	
	[image: image16.png]Yo wl—a




	



[image: image1.png]q—Yawu—la=





[image: image2.png]




[image: image3.png]



MB2005

USER  MANUAL

14643-C
	Indice
Rev.
	Date
Date
	Pages modifiées
Modified pages
	Commentaires
Comments

	First edition-A
	October 2005
	Toutes / All
	Création / First draft

	B
	January 2006
	 p6/p7/p8/p10/p11
	details

	
	
	p12
	MB2005 Noise Figure

	
	
	P9
	Consumption< 300mA

	
	
	P15
	Process of LVDT Adjustment

	
	
	P14/15/15
	Calibration Description

	
	
	P5
	Update of list of calibration and adjustment tools

	C
	January 2007
	P7, P9, P11,  P20
	Supply Voltage / nozzle modification

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	


Ce document ne peut engager MARTEC s'il n'est pas revêtu des signatures des personnes désignées.

This document is not a binding proposal on MARTEC  unless signed by the designated persons.
NdP104-1
Ed 05
	VISA / In charge of
	NOM / NAME
	SIGNATURE
	DATE

	Chef de Projet
Project Manager
	D. PROVENT
	DP
	17/01/2007

	Responsable AQ
QA Manager
	M. MARY
	GV 
po G. VILAIN
	17/01/2007

	Responsable Technique
Technical Manager
	F. RACINE
	FR
	17/01/2007


1. TABLE OF CONTENTS
31.
TABLE OF CONTENTS


42.
List of illustrationS


53.
List of CALIBRATION AND ADJUSTMENT TOOLS


64.
INTRODUCTION


75.
GENERAL DESCRIPTION


75.1.
GENERAL (Figure 3-1)


75.2.
DESCRIPTION (Figure 3-1)


85.3.
SPECIFICATIONS


95.4.
OPERATION PRINCIPLE (Figure 2-2)


95.5.
MB2005 TRANSFERT FUNCTION


136.
OPERATION


136.1.
CONNECTIONS (Figures 3-1, 3-2 and 3-3)


136.2.
FILTERED OUTPUT (Figures 3-3 and figure 3-4)


136.3.
ADJUSTMENT




2. List of illustrationS
FIGURES

Page
2-1
Electronic Noise

11
2-2
Operation Principle - Synoptic

12
3-1
MB2005 Description

16
3-2
Connection between AUBRAC digitizer and MB2005

17
3-3
Comparison Between the DASE and the Expert Group Proposals

18
3-4
Frequency Band

19
3-5
Nozzle dimensions

20
3. List of CALIBRATION AND ADJUSTMENT TOOLS
LIST OF CALIBRATION AND ADJUSTMENT TOOLS
	DESCRIPTION


	SUPPLIER

COMPANY   -   ADDRESS
	FUNCTION

	Low pressure calibrator:
Pressure Manometer and equipment
Air manual pump
Vacuum manual pump

Barometer

Calibrator

	MARTEC

MARTEC

MARTEC

	§ 6. 2.3
§ 6. 2.2
§ 6. 2.4



4. INTRODUCTION
The microbarometer MB2005 was developed in order to detect an air nuclear explosion. The sensitive part is set up with an aneroid barometric bellow which warps under atmospheric pressure change; a LVDT sensor measures this deformation. This one is performant and easy to implement. The electronic noise level is 2 mPa rms, between 1 and 10 Hz.
The filtered output passband is between 0.01 Hz and 27 Hz. It can be modified. A passband of 0.001 Hz to 40 Hz has already been used by the Laboratoire de Géophysique (LDG).
The running control can be ensured by the atmospheric pressure measurement.
Digitization units, developed by LDG, are perfectly designed to use those sensors.

5. GENERAL DESCRIPTION
5.1. GENERAL (Figure 3-1)

The sensitive part is an aneroid barometric bellows (10) made of Durinval. The LVDT (Linear Variable Differential Transformer) displacement sensor (9), joined to a low noise electronics, measures the deformation of the barometric aneroid under atmospheric pressure change. The barometric aneroid displacement sensor unit of each microbarometer is temperature-calibrated to minimise its effects.

5.2. DESCRIPTION (Figure 3-1)

The microbarometer is presented as a cylinder of 15 cm diameter and 32 cm high. Its weight is 7 kg.

A 4-pin plug (5) is used to supply the sensor in 12 Volts. The signals are available on a second 7-pin plug (4).

Pin assignement of plugs
	Power supply plug(5)
	Signal
	Comment
	

	Pin 1 Pin 2 internally connected
	0 V
	
	

	Pin 3 Pin 4 internally connected
	+12 V
	Min 11.2V

Max 13V
	Serial Number NV5xxx

Product xx09415E

Product xx09415F

	
	
	Min 10V

Max 16V
	Serial Number after NV6xxx

Product xx09415 G and more


	Signal output  plug(4)
	Signal
	Comment

	Pin 1
	Filtered OUT-
	20mV/Pa

	Pin 2
	Filtered OUT+
	

	Pin 3
	PA-
	1mV/hPa

	Pin 4
	PA+
	

	Pin 5
	PT9-
	100mV/hPa

	Pin 6
	PT9+
	

	Pin 7
	GND
	


The low part constitutes the measurement chamber (7), it consists of the barometric aneroid (10) and the LVDT displacement sensor (9). The chamber is connected to the atmospheric pressure by four nozzles (6). Each nozzle can receive a microporous hose or other to perform an infrasound filtering network in order to minimise pressure changes due to wind effects and reduce microbarometer background noise.

The high part is watertight and receives two electronic boards: 
· A power supply board (2).

· A measurement board (3).
5.3. SPECIFICATIONS

5.3.1 Microbarometric aneroid characteristics

Provided by LDG, it is made of a metal which does not vary significantly with temperature change.

It is designed to operate between 400 and 1200 hPa.

Means mechanical sensitivity: – 35 nm/Pa.

Acceleration sensitivity: vertical axis: ( 1.5 μm/(m/s2).
5.3.2 Characteristics of the LVDT differential transformer displacement sensor

Provided by LDG.

5.3.3 Characteristics of the measurement electronic board

· Quartz oscillator ultra stable in frequency and voltage.

· LVDT primary excitation frequency: 8 kHz.

· LVDT excitation level: 15.6 V peak to peak

(symetrical excitation).

5.3.4 Characteristics of the microbarometer (Figure 2-1)

The sensor is designed to operate at ( 100 hPa of the atmospheric pressure, (standard 1000 hPa ( 100 hPa). 

Sensitivity before filtering (DC-40 Hz band in PT9) is generally of – 1 mV/Pa, i.e. ( 10 Volts for ( 100 hPa.

Sensitivity after filtering (0.01 - 27 Hz band) is of 20 mV/Pa i.e. ( 10 Volts for ( 5 hPa.

This is the best sensitivity for 20 and 24 bits ADC use.

For 16 bits ADC, we can adjust the sentivity to 100 mV/Pa.

Typical electronic noise in the 0.01-10 Hz band is represented in figure 2.1. 

5.3.5 Environment
Temperature:

Sensors are temperature-calibrated and -compensated in order to reduce their thermal sensitivity. The thermal drift after compensation is less than ( 0.1 hPa/°C.

The temperature range is - 20  to + 60°C (-4 to 140°F)
MB2005 is protected for operation in damp and dusty atmosphere IP 68

MB2005 Complies with the following EMC standards: 
EN61326-1/ EN 55022/ EN61000-4-3/ EN61000-4-6/ EN61000-4-2/ EN61000-4-8

5.3.6 Power Supply

Terminal voltage: 12 Volts
Consumption < 300mA. 
There is no connection between power supply input and output signals.
Recommended Voltage Range: refer to pin assignment table.
5.4. OPERATION PRINCIPLE (Figure 2-2)

An analog oscillator set at 8 kHz is carried out from a quartz and voltage reference oscillator. The voltage hence obtained (TP3-TP4) supplies the LVDT displacement sensor primary. The voltage obtained at secondary is demodulated with analog gates controlled by a signal delivered by the 8 kHz delayed oscillator (PT5). The demodulated signal is then filtered (Low-Pass (LP) 40 Hz filter).

The sensitivity in PT9 is of – 1 mV/Pa (see § 6.3 Adjustment). Before amplification, the DC coupled signal is filtered with a 0.01 Hz high-pass filter. A second output AP measures the Atmospheric Pressure (AP) value and allows the running control of the unit.

[image: image16.png]
5.5. MB2005 TRANSFERT FUNCTION
[image: image17.png]Yo wl—a



[image: image18.png]Yo wl—a







5.5.1. PT9 output

With p=jω

[image: image4.wmf]1

*

1

*

2

*

3

*

4

*

5

*

6

0

*

1

)

(

2

3

4

5

6

+

+

+

+

+

+

+

=

p

a

p

a

p

a

p

a

p

a

p

a

b

p

b

p

H


b1=2.133 E-9
b0=1 E-3a6=1.89 E-17
a6=1.89 E-17
a5=7.03 E-15
a4=3.757 E-11
a3=1.321 E-8
a2=1.823 E-5
a1=5.645 E-3

[image: image5.wmf]1

*

92

.

15

*

077

.

0

*

31

.

318

)

(

2

+

+

=

p

p

p

p

Hs


5.5.2. Amplitude response
[image: image6.png]MB2005 amplitude response, dB, ref. 1 V/Pa

-30

&
<

&
3

&
&

g

- MB2005 theorical ampiitude response +5%
++ MB2005 theorical ampiitude response -5%
— MB2005 amplitude response measured with DASE IS Caliorator

10

10
Frequency, Hz




Electronic noise (band 0.01 - 10 Hz).
Sensitivity : 20 mV/Pa.


[image: image7.png]MB2005 electronic noise PSD, dB, ref. 1 PaffHz

-70

T
— MB2005 electronic noise PSD

10°

Frequency, Hz




FIGURE 2-1 Electronic noise

[image: image8.wmf]LVDT

FILTRAGE

VOLTAGE

REFERENCES

FILTERING

G

QUARTZ

OSCILLATOR

PROM

HIGH-PASS 100 S

-20

AP

1 mV/hPa

S = -1 mV/Pa 

S = 20  mV/Pa

PT1

TP3   TP4

8 KHz

PT6

PT7

PT9

PT12

PT5

-1/100

PT2

ANALOG 

GATES

POW-PASS

40 Hz

RAZ

RELAY


FIGURE 2-2: Synoptic-operation principle
6. OPERATION
6.1. CONNECTIONS (Figures 3-1, 3-2 and 3-3)
The sensor is supplied with + 12 volts by a 4-pins male Jupiter plug (5).

The filtered signal and the atmospheric pressure are connected to the 7-pins female Jupiter plug (4).

The wiring is presented in figures 3-2 with the digitizer unit used at LDG (AUBRAC). In this case, the signal cable is RSN53 and the supply cable is RSN29.  
6.2. FILTERED OUTPUT (Figures 3-3 and figure 3-4)
The characteristics of the standard filtered output are summarised in the table figure 3-4. The characteristics of this output are compared with the Expert Group proposals. It’s a symmetrical output.

6.2.1. Passeband (figure 3-4)

The standard passband is set at 0.01-27 Hz. 

It is possible to modify the cut frequency of the high-pass. You need to contact MARTEC for this operation. 

6.2.2. Sensitivity
The standard sensitivity is of 20 mV/Pa. 

It is possible to change the sensitivity. You need to contact MARTEC for this operation.  

6.3. ADJUSTMENT 

6.3.1. Introduction

In factory:

The LVDT adjustment, the sensitivity adjustment of PT9 filtered and the PA output adjustment are performed in factory. 
At first LVDT adjustment permits, by a mechanical adjustment of LVDT, to centred the PT9 filtered excursion of measurement around the local pressure average. 

At second, the sensitivity of the atmospheric pressure dynamic range is set to +/-100hPa by a potentiometer adjustment.
At third, the Pa output offset for measurement of static output pressure is adjusted to pressure on site, by a potentiometer (1mV/hPa).

On site:

The atmospheric pressure is dependant on altitude. The change is about 100hPa for 1000 meters. So, the LVDT mechanical centring has to be done on site, in order to keep the pressure dynamic range well centred on the site atmospheric pressure average.

The adjustment of sensitivity is not necessary.

The Pa output adjustment has to be done in case of using this output.

6.3.2. LVDT adjustment

Tools used in order to adjust LVDT are multimeter and Allen key.

Operation is the following one:

Calculate U3 : 

U3 (V) = – [AP (hPa) – APA(hPa)] x 0.1.

AP: Atmospheric pressure

APA: Atmospheric pressure average (Aproximative value)

Example: if AP = 890 hPa, U3 = – (890 – 870) x 0.1 = – 2 V.

Connect a voltmeter (on PT9 output) between pins 5 and 6 of the 7 points Jupiter output connector. 
Unscrew the 3 screws.

Then move, up or down the LVDT to obtain U3 +/-0.1V on the multimeter. The voltage has to be checked after the complete screwing of the all 3 screws.

The dynamic pressure range is now well centred on the local atmospheric pressure average.

In factory, the local atmospheric average is fixed to 1000hPa. 

6.3.3. Sensitivity adjustment (figure 2-2)

The sensor maximum dynamic range is set to ( 100 hPa atmospheric pressure. Use a low pressure KELLER calibrator to do calibration.
(1)
Set the sensor to be calibrated + calibrator in high pressure (100 hPa) with air pump. Measure the voltage in PT9. 

Adjust P1 to obtain – 10 Volts. Note the value U1.
(2)

Put the unit in low pressure (– 100 hPa) with vacuum pump. Measure the voltage U2 in PT9 (it must be closed to + 10 Volts). Calculate U’2 = U2 x 20 / ((U2(+(U1().
Adjust P1 to obtain in PT9 U’2.
(3)  
Put the unit in high pressure (100 hPa) with air pump. Measure the voltage U'1 in PT9
(U'2(+(U'1( must be equal to 20 Volts +/- 100mV
Adjust P1 to obtain this result if necessary.

6.3.4. Atmospheric pressure output adjustment

The PA output (between pins 3 and 4 of the 7 pins Jupiter connector) is the conversion of the voltage in mVolt directly to the static atmospheric pressure in hPa.

If the LVDT adjust is necessary, the Pa output is no longer the right value.

So, there are 2 kinds of solution to get the value of atmospheric pressure with MB2005:

First solution: per calculation
Record the local atmospheric pressure LAP at which the LVDT was centred. The right atmospheric pressure can be calculated by adding the difference between the local atmospheric pressure setting LAP and the factory atmospheric pressure FAP (1000hPa), to the Pa output value in mV.

Right value = PA + (LAP-FAP)   each value in hPa.

Second solution: per adjustment
This adjustment needs an external reference barometer. It is necessary to adjust the value of the potentiometer P3 on the measurement board by the use of a small screwdriver.

Turn the potentiometer until the voltage value (mVolt) of the Pa output equal the local atmospheric pressure (hPa).

[image: image9.png]10.

e E NP B oo

COVER.

POWER SUPPLY BOARD.
MEASUREMENT BOARD.
SIGNAL OUTPUT PLUG.
POWER SUPPLY PLUG.
NOZZLES (4).

MEASUREMENT CHAMBER.
LVDT JUNGTION CONNECTOR.
LVDT DISPLACEMENT SENSOR.
BAROMETRIC ANEROID.





FIGURE 3-1: MB2005 Description

[image: image10.png]JUPITER AUBRAC TYPE UNIT
MEASUREMENT 10TP7 JUPITER Fe=12elh

BOARD FILECA CABLE PART N°4144/3 107P5 8 bits

| 4 JAimospheric pressure OR 2PAIRS +1 SHIELDED
| CABLE

CONTROL CP BOARD

Fe=1t0200e/s
20 or 24 bits
]

P9 2 MEASUREMENT CABLE \_.

1 DIFFERENTIAL AMPLIFIERS

Filtered ouput DIGITIZER BOARDS
—® [ 4
5
L 2
SENSOR MECHANICAL GROUND
CONNECTION OF
*5v 15V 0 - 18V SHIELDING BRAID
PWR SUPP BJARD ON CONNECTOR
L 2 L 4 &
CONVERTERS °
:]O_: 12V [ 4/~ 15V .
g .
D
[G] ® & 9
i
'
12 VOLTS POWER SUPPLY CABLE 12v
4
L L ]
3 ( ) .
2 @ ®
J
1
SENSOR

JUPITER 10TP4





FIGURE 3-2: Connection between AUBRAC digitizer and MB2005 
	
	
	MB2005

Infrasonic

Sensor
	Specifications

CTBT/PC/II/WGB/1

	
	Range
	200 hPa pp
	

	
	Bandwidth (Hz)
	0 – 40
	

	
	Sensitivity
	1mV / Pa
	

	
	Electronic noise (at 1 Hz)
	0.6 mPa/
[image: image11.wmf]Hz


	
[image: image12.wmf]£

 0.63 mPa / 
[image: image13.wmf]Hz



	
	Electronic noise (0.02-4Hz)
	2 mPa rms
	

	
	Dynamic range (1)
	134 dB
	

	
	Mode/Type
	Single Ended
	

	
	Range
	1000 Pa pp
	

	
	Bandwidth (Hz)
	0.01 – 27
	0.02-4Hz

	
	Sensitivity
	20 mV / Pa
	

	
	Electronic noise (0.02-4Hz)
	2 mPa rms
	

	
	Dynamic range (1)
	108 dB
	=108dB

	
	Mode/Type
	Differential
	

	
	Range
	200 Pa pp

(within 500 to 1200hPa)
	

	
	Sensitivity
	1 mV / hPa
	


(1) ½ Full scale/Noise.

[image: image14.png]0ol

(zH) Asuanbaiy

0ol 10 100 L00'0
00'0¢-
- —+ - e - - 00'02-
ZH 20°0 \
\\ /100 |
/ /
/1
\ \
/
/
- //I.IIJ. —I- miut b e e et .-I\N\JI-%%!I%IIJI 00°€-
/; \\\\\
= 000

Attenuation (dB)





FIGURE 3-4: Frequency Band
[image: image15.emf]
FIGURE 3-5: Nozzle dimensions
(applied since ref. xx09451 F – except serial number NV5166 / NV5167 / NV5169 and NV5170)



MB2005 �Unfilterd� output





MB2005 �Filtered  �output 





MB2005 PA output





Pt9 output





Filtered output





Pressure





    Hs(p)





    H(p)








	MB2005 USER MANUAL
Reference : 14643-C
	15/20
	Edition : 17/01/2007



_1186236622.bin

_1190468994.unknown

_1205832042.unknown

_1190029277.unknown

_1190029278.unknown

_1190029276.unknown

_1186234923.bin

