## Slide 1: New Optical Microbarometer
26th june 2017
SnT 2017 | Serge Olivier (CEA); Anthony Hue, Nathalie Olivier, Serge Le Mallet (SeismoWave)
This project is funded by the DGA though convention n°132906059

## Slide 2: SUMMARY
  INTRODUCTION
  MICROBAROMETER PRINCIPLES
  OPTICAL TRANSDUCER PRINCIPLES
  OPTICAL MICROBAROMETER
  THE FIRST DESIGN
  THE SECOND DESIGN
  CONCLUSION

## Slide 3: INTRODUCTION
  Etc…
  Aneroid bellows
CEA microbarometer designs:
… 2000 2005 2015 Today -> Tomorrow
  MB2000 MB2005 MB3
  Interferometer
  LVDT transducer
  Magnet & coil
  transducer
  R&D : OptoGeo new optical microbarometer
  French Patent 1262878
R&D
  Seismic pendulum
  Movement transducers
  Digitizing systems
  Metrology process
  Electronic devices

## Slide 4: INTRODUCTION
  Principles
  Interferometer compared to LVDT and Magnet and Coil transducers:
      Lower noise
      Larger operational range
        Two quadrature fringes
        No limit on phase fringes measurement
  Objectives
  Increase performances (noise and operational range) of the sensor
  No more electrical mass difficulty

## Slide 5: MICROBAROMETER PRINCIPLES
Current microbarometer principle
dP +
dP -
P +
P -
SIGNAL AMPLIFICATION
SIGNAL SYMMETRIZATION
ACTIVE LOW PASS FILTER
SIGNAL SYMMETRIZATION
COIL OUTPUT VOLTAGE
Pressure Dérivative
Pressure
DIGITIZING UNIT
Magnet & coil transducer
Aneroid bellows
Measurement cavity
MB3a Hood
Transducer block
ATMOSPHERIC PRESSURE
MB3a microbarometer

## Slide 6: MICROBAROMETER PRINCIPLES
dP +
dP -
P +
P -
SIGNAL AMPLIFICATION
SIGNAL SYMMETRIZATION
ACTIVE LOW PASS FILTER
SIGNAL SYMMETRIZATION
COIL OUTPUT VOLTAGE
Pressure Dérivative
Pressure
DIGITIZING UNIT
Magnet & coil transducer
Aneroid bellows
Measurement cavity
MB3a Hood
Transducer block
ATMOSPHERIC PRESSURE
MB3a microbarometer
Magnet and coil transducers principle
Replaced by an optical transducer and a specific digitizer
OK

## Slide 7: Optical microbarometer
Optical signals
- DIGITIZING and PROCESSING
System
- laser source
Moving miror
Aneroid bellows
Measurement cavity
Interferometer
Transducer block
ATMOSPHERIC PRESSURE
Aneroid
bellows
Interferometer
Optical microbarometer
Moving mirror (Cat’s eye)
Pa/s / count
Bellows sensitivity = 122nm/Pa

## Slide 8: Interferometric Transducer
  Interferometer is designed with optical waveguide obtained by ion-exchange process into a glass wafer
  Less adjustment than standard interferometer
  Interferometer is included into a compact package
Interferometer principle
Interferometer technology: integrated optics
  « Mic-Mac » interferometer: One Michelson arm and one Mach-Zehnder arm
  2 quadrature fringes to compute the phase:
      Meas1 & Meas2
  2 photometric signals to control photometric variations
      Ref1 & Ref2

## Slide 9: Optical microbarometer
  Step 1: Concept validation
    Interferometer outside the aneroid bellows
  Step 2: Fully integrated prototype
    Interferometer inside the aneroid bellows
2 steps to design the optical microbarometer

## Slide 10: First design
6 prototypes
  3 with a MB3a bellows (OptoGeo 1,2,3)
  3 with a MB2005 aneroid bellows (OptoGeo 4,5,6)
Objectives:
  to characterize the main components (aneroid bellows, interferometer)

## Slide 11: First design measurements
Results:
  Signal PSD is similar for all sensors in a wide bandwidth
  High frequency: seismic sensitivity
Experiment:
  6 OptoGeo + 1 MB2005 (reference sensor)
  Observations of the same signal
  Comparison between each sensor measurement to validate the OptoGeo measurement

## Slide 12: First design measurements
Coherence:
  Between MB2005 and each OptoGeo:
      > 0,98 on a large bandwidth
      f > 1 Hz: noise level of MB2005
  Between 2 OptoGeo:
      > 0,98 on all the bandwidth
Noise evaluation:
  2 sensors method (Holcomb, L.G., 1989)
  Better noise for OptoGeo than MB2005
  Between 0,1 and 0,3 Hz: due to evaluation method
  Optimization of measurement cavity
  Noise measurement
  Better noise than MB2005 and MB3

## Slide 13: First design measurements
Thermal test:
  Temperature from 5 to 40 °C
  Steps = 5 °C

## Slide 14: First design measurements
Results:
  Thermal effects on sensitivity comparable to MB2005

## Slide 15: First design measurements
Effects of altitude:
  Expected: no effect due to interferometric measurement principle
  Transversal deformation due to symmetrical errors during manufacturing
  Measurement: The 2 photometric signals at 5 different altitudes
Results:
  No complete loss of reflected optical beam: OptoGeo always gives reliable measurements
  Lower effects for OptoGeo with MB3 aneroid bellows
  Larger operational range than MB3 & MB2005
  Optogeo will be usable in all CTBTO stations without any adjustment

## Slide 16: First design
Sum up:
  Significantly lower noise than MB2005 and slightly lower than MB3
  Temperature effects on sensitivity: similar to MB2005/MB3 microbarometers
  Altitude effects:
      Larger operational range than MB3&MB2005
      No adjustment
Conclusion:
  Encouraging results for the following design
  Choice of MB3 aneroid bellows
  Following design: interferometer inside the aneroid bellows
      optical components protection (humidity, dust, …)

## Slide 17: Second design
Adjustment system
Vacuum
valve
interferometer
sealed leading-through device
Goal:
  Insert the interferometer inside the aneroid bellows, under vacuum to avoid environmental perturbations (humidity, dust, …)
4 prototypes
  With MB3 aneroid bellows
  Challenges:
      Smaller interferometer under vacuum
      To adjust the interferometer and mirror positions from the outside of aneroid bellows

## Slide 18: Second design experiment
Effects of the test bench
  Flat response on a large bandwidth
  +/- 0,5 dB
Transfer function measurement:
  Use of infrason metrology test bench (IS Calibrator)

## Slide 19: Conclusion
  Robustness of optical fiber connectors
  Mechanical design to improve (adjustment bellows, …)
  Lower noise than MB2005 on all the bandwidth
  Temperature effects: similar to present microbarometers
  Altitude effects: no adjustment
      Better static dynamic
  Flat instrumental response on all the bandwidth: interesting for metrology (possibility for secondary standard)
  Very interesting results
  Feasibility prouved
SUM UP
FUTURE WORKS
  « Optical » digitizer: lower noise, lower consumption
  Quality of the laser source (lower phase noise)
  Application for seismic VBB sensors: launching a new 3 years program of R&D (Project under contractualization with DGA)

## Slide 20: DAM DASE STMG
Commissariat à l’énergie atomique et aux énergies alternatives
Centre DIF | 91297 Arpajon Cedex
  Etablissement public à caractère industriel et commercial | RCS Paris B 775 685 019