## Slide 1
S. OLIVIER CEA DAM DIF, F-91297
F-91297 Arpajon, France, serge.olivier@cea.fr
Devices under test:
One microbarometer MB2005
Six optical microbarometers:
Optogeo 1, Optogeo 2 and Optogeo 3 used a MB3
aneroid capsule with the corresponding sensitivity
Optogeo 4, Optogeo 5 and Optogeo 6 used a MB2005
aneroid capsule with the corresponding sensitivity
PSD & coherency:
good agreement. Optogeo sensitivity is OK
Noise evaluation:
better than MB2005 for all the bandwidth
Comparison with MB2005
First prototype
Optical area
Interferometer
Fiber
optic
Aneroid
capsule
2 prototypes
Digital microbarometer
Analog microbarometer
Digitizer
Mechanical transducer
Movement transducer
Pressure
variation
Displacement
Analog voltage
Digital data
Current microbarometer
Optical microbarometer
Microbarometer principle
Aneroid
capsule
LVDT or
Magnet-coil
1
2
3
project funded by
A. HUE, N. OLIVIER, S. LE MALLET
SEISMO WAVE, Lannion, France,
www.seismowave.com
Effects of temperature
4
Axial deformation: measurement principle
Transversal deformation: due to symmetrical errors during manufacturing
Risk: misalignment due to transversal deformation
Measurement: Observation of Optical beam intensity variation with altitude
Effects of altitude
Final tests
5
6
Moving mirror
Aneroid
capsule
Alt.= 0 m
Aneroid capsule
Alt.= 2000 m
Optical beam
Moving mirror
Aneroid
capsule
Vacuum
Optical beam
Interferometer
Very low effects at an altitude of 2000 m for sensor with MB3 aneroid capsule
Low effects at an altitude of 2000 m for sensor with MB2005 aneroid sensor
Low thermal effects on infrasound sensitivity compared to MB2005*
* MB2005 thermal drift is less than ±0.1 hPa/°C
Thermal test:
From 5 °C to 40 °C
Step : 5 °C
New optical microbarometer
EGU 2017
X3.27 7505
through convention n°132906059
Test of 2 types of capsules
MB2000 and MB3 in ambient environment
1 MB3 bellow and interferometer under vacuum
Aneroid capsule area
second prototype
The OptoGeo self noise (Green) is bellow than the MB3a self noise (blue) and follow the digitizer noise level (Cyan) which limits the measurement
Flat response of the transfer function on the bandwidth 0,001 Hz à 50 Hz
The phase of the sensor has a constant value unlike the other microbarometers used.
Moving
mirror
Optical
beam
Moving mirror
Vacuum
Optical beam
Interferometer
Vacuum
Metrology Test Bench
Flat response of the transfer function: the defaults observed come from the environmental effects during the measurements in LF and in HF from the calibration speaker resonance
Conclusion
7
Advantages:
Better resolution for the total bandwidth
Less sensitivity to temperature variation
No adjustment with altitude
Less optical adjustment compared with other optical technology
Other interests:
OptoGeo sensor is a useful tool to characterize metrology test bench on a wider frequency bandwidth
Future works:
• Optical digitizer: lower noise, lower consumption
• Quality of the laser source
• Calibration capability
• Low frequencies test bench