## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 280 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 245
Hood | 02-92B/P | 201
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1007.7 | hPa |  | PRC PRO 01
Humidity | 51.5 | %/Hr
Internal temperature of calibration bench | 23.94 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1014 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.4 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.93 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.99 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3094 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 17/09/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 18/09/2018
Technical Manager | Anthony HUE | Approval |  | 27/09/2018
CEO | Robert GLEMOT | Approval |  | 27/09/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.5166 | 57.6398
0.012209 | -36.0789 | 38.5848
0.018313 | -35.0249 | 28.3574
0.024417 | -34.5614 | 22.0932
0.030521 | -34.4345 | 18.2959
0.036626 | -34.291 | 15.2381
0.04273 | -34.1969 | 12.9014
0.048834 | -34.1498 | 11.3553
0.054938 | -34.0877 | 10.207
0.061043 | -34.029 | 9.1639
0.067147 | -34.0164 | 8.3483
0.073251 | -34.0163 | 7.5779
0.079355 | -34.0331 | 6.8957
0.08546 | -34.0444 | 6.2902
0.091564 | -34.0357 | 5.8446
0.097668 | -34.014 | 5.5838
0.10377 | -33.9909 | 5.32
0.10988 | -33.9728 | 5.0277
0.11598 | -33.9636 | 4.7236
0.12209 | -33.9656 | 4.4347
0.12819 | -33.9756 | 4.1575
0.13429 | -33.9858 | 3.8834
0.1404 | -33.9878 | 3.6171
0.1465 | -33.9889 | 3.3701
0.15261 | -33.9859 | 3.1612
0.15871 | -33.9799 | 3.0105
0.16482 | -33.9728 | 2.9182
0.17092 | -33.9706 | 2.8831
0.17702 | -33.9688 | 2.8145
0.18313 | -33.9682 | 2.7344
0.18923 | -33.9686 | 2.6216
0.19534 | -33.9686 | 2.487
0.20144 | -33.9688 | 2.3361
0.20754 | -33.967 | 2.2098
0.21365 | -33.9669 | 2.0901
0.21975 | -33.9667 | 1.9893
0.22586 | -33.9676 | 1.9022
0.23196 | -33.9695 | 1.831
0.23807 | -33.9727 | 1.7709
0.24417 | -33.976 | 1.7169
0.25027 | -33.9794 | 1.6587
0.25638 | -33.9814 | 1.6022
0.26248 | -33.983 | 1.5416
0.26859 | -33.9839 | 1.4822
0.27469 | -33.9843 | 1.4165
0.2808 | -33.9852 | 1.3445
0.2869 | -33.9859 | 1.2751
0.293 | -33.9866 | 1.2083
0.29911 | -33.9873 | 1.1465
0.30521 | -33.9879 | 1.0848
0.31132 | -33.9883 | 1.0405
0.31742 | -33.9892 | 0.99342
0.32353 | -33.9893 | 0.95369
0.32963 | -33.99 | 0.91934
0.33573 | -33.9907 | 0.89311
0.34184 | -33.9917 | 0.86643
0.34794 | -33.9935 | 0.83928
0.35405 | -33.9952 | 0.8112
0.36015 | -33.9973 | 0.78279
0.36626 | -33.999 | 0.75241
0.37236 | -34.0008 | 0.71433
0.37846 | -34.0026 | 0.66933
0.38457 | -34.0034 | 0.61345
0.39067 | -34.0043 | 0.56683
0.39678 | -34.0049 | 0.51587
0.40288 | -34.0054 | 0.46317
0.40899 | -34.0053 | 0.41254
0.41509 | -34.0048 | 0.36284
0.42119 | -34.0049 | 0.32097
0.4273 | -34.0049 | 0.27611
0.4334 | -34.005 | 0.23117
0.43951 | -34.0054 | 0.18913
0.44561 | -34.0062 | 0.15015
0.45172 | -34.0076 | 0.11824
0.45782 | -34.0083 | 0.081639
0.46392 | -34.0094 | 0.045906
0.47003 | -34.0102 | 0.013058
0.47613 | -34.0109 | -0.020871
0.48224 | -34.0113 | -0.055536
0.48834 | -34.0103 | -0.089841
0.49445 | -34.0099 | -0.1246
0.50055 | -34.0094 | -0.15622
0.50665 | -34.0086 | -0.18966
0.51276 | -34.0076 | -0.22261
0.51886 | -34.0067 | -0.25524
0.52497 | -34.0056 | -0.27804
0.53107 | -34.0048 | -0.30611
0.53717 | -34.0042 | -0.33279
0.54328 | -34.0035 | -0.35436
0.54938 | -34.0033 | -0.37558
0.55549 | -34.0034 | -0.39519
0.56159 | -34.0043 | -0.40608
0.5677 | -34.0045 | -0.41989
0.5738 | -34.0054 | -0.43514
0.5799 | -34.0061 | -0.4492
0.58601 | -34.007 | -0.45892
0.59211 | -34.0083 | -0.46548
0.59822 | -34.0095 | -0.47483
0.60432 | -34.0107 | -0.48644
0.61043 | -34.0117 | -0.4946
0.61653 | -34.0126 | -0.50485
0.62263 | -34.0139 | -0.51439
0.62874 | -34.0143 | -0.52552
0.63484 | -34.0149 | -0.53504
0.64095 | -34.0152 | -0.54567
0.64705 | -34.0154 | -0.55587
0.65316 | -34.0155 | -0.56861
0.65926 | -34.0155 | -0.58454
0.66536 | -34.0144 | -0.60214
0.67147 | -34.014 | -0.61577
0.67757 | -34.0136 | -0.63042
0.68368 | -34.0129 | -0.64914
0.68978 | -34.0122 | -0.66643
0.69589 | -34.0113 | -0.68541
0.70199 | -34.0105 | -0.70775
0.70809 | -34.0097 | -0.72686
0.7142 | -34.0091 | -0.74719
0.7203 | -34.0081 | -0.76635
0.72641 | -34.0075 | -0.78852
0.73251 | -34.007 | -0.80845
0.73862 | -34.0072 | -0.82994
0.74472 | -34.0069 | -0.84991
0.75082 | -34.0067 | -0.87178
0.75693 | -34.0071 | -0.89341
0.76303 | -34.0076 | -0.91505
0.76914 | -34.0093 | -0.93524
0.77524 | -34.0102 | -0.95481
0.78135 | -34.0114 | -0.97476
0.78745 | -34.0127 | -0.99463
0.79355 | -34.0139 | -1.0143
0.79966 | -34.0155 | -1.0336
0.80576 | -34.0176 | -1.0533
0.81187 | -34.0191 | -1.0746
0.81797 | -34.0209 | -1.0956
0.82408 | -34.0225 | -1.116
0.83018 | -34.0241 | -1.1364
0.83628 | -34.0258 | -1.1549
0.84239 | -34.0269 | -1.175
0.84849 | -34.0282 | -1.1963
0.8546 | -34.0296 | -1.2159
0.8607 | -34.0307 | -1.2362
0.86681 | -34.0317 | -1.2549
0.87291 | -34.0326 | -1.2735
0.87901 | -34.033 | -1.2911
0.88512 | -34.0336 | -1.3091
0.89122 | -34.0343 | -1.3271
0.89733 | -34.0346 | -1.3426
0.90343 | -34.035 | -1.36
0.90953 | -34.0351 | -1.3743
0.91564 | -34.035 | -1.39
0.92174 | -34.0352 | -1.4049
0.92785 | -34.0352 | -1.4194
0.93395 | -34.0352 | -1.4332
0.94006 | -34.0353 | -1.4462
0.94616 | -34.035 | -1.4582
0.95226 | -34.035 | -1.471
0.95837 | -34.035 | -1.4847
0.96447 | -34.0351 | -1.4975
0.97058 | -34.0349 | -1.5109
0.97668 | -34.0346 | -1.5243
0.98279 | -34.0342 | -1.5371
0.98889 | -34.0339 | -1.548
0.99499 | -34.0338 | -1.5615
1.0011 | -34.0336 | -1.5724
1.0072 | -34.0332 | -1.5842
1.0133 | -34.0328 | -1.5953
1.0194 | -34.0324 | -1.6083
1.0255 | -34.032 | -1.622
1.0316 | -34.0314 | -1.6341
1.0377 | -34.031 | -1.6477
1.0438 | -34.0304 | -1.6605
1.0499 | -34.0297 | -1.6756
1.056 | -34.029 | -1.6898
1.0621 | -34.0284 | -1.704
1.0682 | -34.0279 | -1.7165
1.0743 | -34.0273 | -1.7317
1.0805 | -34.0268 | -1.7459
1.0866 | -34.0259 | -1.7606
1.0927 | -34.0251 | -1.775
1.0988 | -34.0246 | -1.7896
1.1049 | -34.0238 | -1.8037
1.111 | -34.0231 | -1.8187
1.1171 | -34.0225 | -1.833
1.1232 | -34.0221 | -1.8478
1.1293 | -34.0217 | -1.8616
1.1354 | -34.0214 | -1.8763
1.1415 | -34.0211 | -1.8914
1.1476 | -34.0206 | -1.9065
1.1537 | -34.0201 | -1.9208
1.1598 | -34.0203 | -1.9352
1.1659 | -34.0201 | -1.9483
1.172 | -34.0199 | -1.962
1.1781 | -34.0197 | -1.9757
1.1842 | -34.0196 | -1.9896
1.1903 | -34.0199 | -2.0006
1.1964 | -34.0198 | -2.0139
1.2025 | -34.0199 | -2.0276
1.2086 | -34.0202 | -2.04
1.2147 | -34.0203 | -2.0524
1.2209 | -34.0206 | -2.0637
1.227 | -34.0212 | -2.0747
1.2331 | -34.0216 | -2.0867
1.2392 | -34.0219 | -2.0985
1.2453 | -34.0223 | -2.1098
1.2514 | -34.0227 | -2.1199
1.2575 | -34.0232 | -2.131
1.2636 | -34.0239 | -2.1404
1.2697 | -34.0245 | -2.1499
1.2758 | -34.0251 | -2.1612
1.2819 | -34.0257 | -2.1714
1.288 | -34.0263 | -2.1824
1.2941 | -34.027 | -2.1913
1.3002 | -34.0276 | -2.1999
1.3063 | -34.0281 | -2.2094
1.3124 | -34.0286 | -2.2189
1.3185 | -34.0293 | -2.2283
1.3246 | -34.0299 | -2.2379
1.3307 | -34.0306 | -2.2469
1.3368 | -34.0313 | -2.2562
1.3429 | -34.0319 | -2.2655
1.349 | -34.0323 | -2.2753
1.3551 | -34.0329 | -2.285
1.3613 | -34.0334 | -2.2937
1.3674 | -34.0339 | -2.3042
1.3735 | -34.0343 | -2.3142
1.3796 | -34.0347 | -2.3238
1.3857 | -34.035 | -2.3335
1.3918 | -34.0355 | -2.344
1.3979 | -34.036 | -2.3535
1.404 | -34.0364 | -2.3651
1.4101 | -34.0367 | -2.3757
1.4162 | -34.0371 | -2.3865
1.4223 | -34.0375 | -2.3981
1.4284 | -34.0379 | -2.4083
1.4345 | -34.0381 | -2.42
1.4406 | -34.0384 | -2.4334
1.4467 | -34.0388 | -2.4451
1.4528 | -34.0392 | -2.4566
1.4589 | -34.0396 | -2.4685
1.465 | -34.0398 | -2.4796
1.4711 | -34.0403 | -2.4911
1.4772 | -34.0407 | -2.5042
1.4833 | -34.0411 | -2.5172
1.4894 | -34.0416 | -2.5294
1.4955 | -34.0421 | -2.5414
1.5016 | -34.0426 | -2.5538
1.5078 | -34.0429 | -2.5675
1.5139 | -34.0433 | -2.5804
1.52 | -34.0437 | -2.593
1.5261 | -34.0439 | -2.6055
1.5322 | -34.0442 | -2.6193
1.5383 | -34.0446 | -2.6318
1.5444 | -34.0451 | -2.6444
1.5505 | -34.0455 | -2.6576
1.5566 | -34.0459 | -2.6705
1.5627 | -34.0464 | -2.6841
1.5688 | -34.0469 | -2.6975
1.5749 | -34.0472 | -2.7101
1.581 | -34.0475 | -2.7233
1.5871 | -34.0478 | -2.7369
1.5932 | -34.0481 | -2.7496
1.5993 | -34.0485 | -2.7636
1.6054 | -34.0488 | -2.7767
1.6115 | -34.0492 | -2.7906
1.6176 | -34.0495 | -2.8031
1.6237 | -34.0496 | -2.8166
1.6298 | -34.0499 | -2.8292
1.6359 | -34.0502 | -2.8419
1.642 | -34.0506 | -2.8539
1.6482 | -34.0507 | -2.867
1.6543 | -34.051 | -2.8805
1.6604 | -34.0513 | -2.8925
1.6665 | -34.0514 | -2.9049
1.6726 | -34.0515 | -2.9175
1.6787 | -34.0518 | -2.9301
1.6848 | -34.0518 | -2.9427
1.6909 | -34.052 | -2.9555
1.697 | -34.0522 | -2.9677
1.7031 | -34.0522 | -2.9804
1.7092 | -34.0522 | -2.9926
1.7153 | -34.0523 | -3.0047
1.7214 | -34.0521 | -3.0169
1.7275 | -34.0521 | -3.0296
1.7336 | -34.052 | -3.0425
1.7397 | -34.0521 | -3.0548
1.7458 | -34.0521 | -3.0671
1.7519 | -34.0521 | -3.0793
1.758 | -34.0522 | -3.0914
1.7641 | -34.0523 | -3.103
1.7702 | -34.0522 | -3.116
1.7763 | -34.0522 | -3.1276
1.7824 | -34.0522 | -3.1401
1.7885 | -34.0522 | -3.1525
1.7947 | -34.0523 | -3.1646
1.8008 | -34.0523 | -3.176
1.8069 | -34.0523 | -3.1877
1.813 | -34.0523 | -3.1999
1.8191 | -34.0523 | -3.2126
1.8252 | -34.0524 | -3.2253
1.8313 | -34.0525 | -3.2375
1.8374 | -34.0527 | -3.2494
1.8435 | -34.0527 | -3.2613
1.8496 | -34.0528 | -3.2742
1.8557 | -34.0529 | -3.2869
1.8618 | -34.0532 | -3.2989
1.8679 | -34.0533 | -3.3109
1.874 | -34.0534 | -3.3228
1.8801 | -34.0536 | -3.3351
1.8862 | -34.0538 | -3.3471
1.8923 | -34.0539 | -3.3591
1.8984 | -34.0541 | -3.371
1.9045 | -34.0542 | -3.3825
1.9106 | -34.0542 | -3.3946
1.9167 | -34.0543 | -3.4066
1.9228 | -34.0544 | -3.4184
1.9289 | -34.0545 | -3.4301
1.9351 | -34.0547 | -3.4409
1.9412 | -34.0549 | -3.4521
1.9473 | -34.0552 | -3.464
1.9534 | -34.0554 | -3.4753
1.9595 | -34.0556 | -3.4869
1.9656 | -34.0558 | -3.4989
1.9717 | -34.056 | -3.5105
1.9778 | -34.0563 | -3.5218
1.9839 | -34.0566 | -3.5331
1.99 | -34.0567 | -3.5441
1.9961 | -34.0568 | -3.5561
2.0022 | -34.057 | -3.5673
2.0083 | -34.0571 | -3.5784
2.0144 | -34.0573 | -3.5897
2.0205 | -34.0576 | -3.6013
2.0266 | -34.0579 | -3.6125
2.0327 | -34.0582 | -3.6239
2.0388 | -34.0585 | -3.6352
2.0449 | -34.0587 | -3.6465
2.051 | -34.059 | -3.6577
2.0571 | -34.0593 | -3.6689
2.0632 | -34.0595 | -3.6802
2.0693 | -34.0596 | -3.6921
2.0754 | -34.0599 | -3.7033
2.0816 | -34.0601 | -3.7151
2.0877 | -34.0603 | -3.7265
2.0938 | -34.0606 | -3.7384
2.0999 | -34.0608 | -3.7499
2.106 | -34.061 | -3.7613
2.1121 | -34.0611 | -3.7731
2.1182 | -34.0613 | -3.7837
2.1243 | -34.0616 | -3.7943
2.1304 | -34.0617 | -3.8053
2.1365 | -34.062 | -3.8169
2.1426 | -34.0622 | -3.8286
2.1487 | -34.0625 | -3.8397
2.1548 | -34.0628 | -3.8513
2.1609 | -34.0632 | -3.8626
2.167 | -34.0635 | -3.8736
2.1731 | -34.0638 | -3.8854
2.1792 | -34.064 | -3.8963
2.1853 | -34.0643 | -3.908
2.1914 | -34.0646 | -3.9192
2.1975 | -34.0649 | -3.9305
2.2036 | -34.0651 | -3.9424
2.2097 | -34.0653 | -3.9536
2.2158 | -34.0656 | -3.9642
2.222 | -34.0659 | -3.9746
2.2281 | -34.0661 | -3.9858
2.2342 | -34.0663 | -3.997
2.2403 | -34.0665 | -4.0075
2.2464 | -34.0667 | -4.0182
2.2525 | -34.0669 | -4.0295
2.2586 | -34.0672 | -4.04
2.2647 | -34.0675 | -4.0504
2.2708 | -34.0678 | -4.0615
2.2769 | -34.068 | -4.0721
2.283 | -34.0682 | -4.0831
2.2891 | -34.0684 | -4.0939
2.2952 | -34.0685 | -4.1044
2.3013 | -34.0687 | -4.1152
2.3074 | -34.0689 | -4.1256
2.3135 | -34.0691 | -4.1369
2.3196 | -34.0693 | -4.147
2.3257 | -34.0695 | -4.1576
2.3318 | -34.0696 | -4.1685
2.3379 | -34.0697 | -4.1793
2.344 | -34.0699 | -4.1905
2.3501 | -34.07 | -4.2011
2.3562 | -34.0702 | -4.2116
2.3623 | -34.0703 | -4.2219
2.3685 | -34.0704 | -4.2324
2.3746 | -34.0705 | -4.2436
2.3807 | -34.0707 | -4.2541
2.3868 | -34.0708 | -4.2648
2.3929 | -34.071 | -4.2754
2.399 | -34.071 | -4.2858
2.4051 | -34.0713 | -4.2963
2.4112 | -34.0714 | -4.3069
2.4173 | -34.0716 | -4.3175
2.4234 | -34.0716 | -4.3283
2.4295 | -34.0718 | -4.3391
2.4356 | -34.0719 | -4.3499
2.4417 | -34.072 | -4.361
2.4478 | -34.0721 | -4.3719
2.4539 | -34.0722 | -4.3821
2.46 | -34.0723 | -4.3935
2.4661 | -34.0725 | -4.4046
2.4722 | -34.0726 | -4.4148
2.4783 | -34.0727 | -4.4253
2.4844 | -34.0727 | -4.4363
2.4905 | -34.0728 | -4.4473
2.4966 | -34.073 | -4.458
2.5027 | -34.0731 | -4.4685
2.5089 | -34.0732 | -4.4794
2.515 | -34.0732 | -4.4902
2.5211 | -34.0733 | -4.501
2.5272 | -34.0734 | -4.512
2.5333 | -34.0735 | -4.5229
2.5394 | -34.0736 | -4.5337
2.5455 | -34.0737 | -4.545
2.5516 | -34.0739 | -4.5554
2.5577 | -34.0741 | -4.5668
2.5638 | -34.0742 | -4.5785
2.5699 | -34.0742 | -4.5899
2.576 | -34.0743 | -4.6014
2.5821 | -34.0744 | -4.6127
2.5882 | -34.0745 | -4.6238
2.5943 | -34.0746 | -4.6346
2.6004 | -34.0748 | -4.6456
2.6065 | -34.0749 | -4.6561
2.6126 | -34.075 | -4.6667
2.6187 | -34.0751 | -4.6773
2.6248 | -34.0752 | -4.6881
2.6309 | -34.0754 | -4.6995
2.637 | -34.0756 | -4.7111
2.6431 | -34.0758 | -4.7217
2.6492 | -34.0759 | -4.7331
2.6554 | -34.0761 | -4.744
2.6615 | -34.0762 | -4.755
2.6676 | -34.0764 | -4.7664
2.6737 | -34.0766 | -4.7775
2.6798 | -34.0769 | -4.7887
2.6859 | -34.0771 | -4.8
2.692 | -34.0772 | -4.8109
2.6981 | -34.0774 | -4.822
2.7042 | -34.0776 | -4.8332
2.7103 | -34.0778 | -4.8439
2.7164 | -34.0781 | -4.8549
2.7225 | -34.0783 | -4.8662
2.7286 | -34.0785 | -4.8773
2.7347 | -34.0787 | -4.8879
2.7408 | -34.079 | -4.8981
2.7469 | -34.0792 | -4.9094
2.753 | -34.0794 | -4.92
2.7591 | -34.0797 | -4.9313
2.7652 | -34.0799 | -4.9428
2.7713 | -34.08 | -4.9536
2.7774 | -34.0803 | -4.9646
2.7835 | -34.0805 | -4.9756
2.7896 | -34.0807 | -4.9861
2.7958 | -34.0809 | -4.997
2.8019 | -34.0812 | -5.0074
2.808 | -34.0814 | -5.0183
2.8141 | -34.0817 | -5.0294
2.8202 | -34.082 | -5.0403
2.8263 | -34.0823 | -5.0512
2.8324 | -34.0826 | -5.0626
2.8385 | -34.083 | -5.0734
2.8446 | -34.0833 | -5.0842
2.8507 | -34.0836 | -5.095
2.8568 | -34.084 | -5.1056
2.8629 | -34.0842 | -5.1162
2.869 | -34.0845 | -5.1273
2.8751 | -34.0848 | -5.138
2.8812 | -34.0852 | -5.1492
2.8873 | -34.0854 | -5.1602
2.8934 | -34.0857 | -5.1712
2.8995 | -34.086 | -5.1822
2.9056 | -34.0863 | -5.1933
2.9117 | -34.0866 | -5.2043
2.9178 | -34.0869 | -5.2151
2.9239 | -34.0873 | -5.2265
2.93 | -34.0877 | -5.2379
2.9361 | -34.0879 | -5.2487
2.9423 | -34.0882 | -5.2594
2.9484 | -34.0886 | -5.2704
2.9545 | -34.089 | -5.2813
2.9606 | -34.0893 | -5.2926
2.9667 | -34.0897 | -5.3037
2.9728 | -34.09 | -5.3149
2.9789 | -34.0904 | -5.3259
2.985 | -34.0907 | -5.3371
2.9911 | -34.0911 | -5.3482
2.9972 | -34.0915 | -5.3593
3.0033 | -34.0918 | -5.3707
3.0094 | -34.0922 | -5.3817
3.0155 | -34.0926 | -5.3925
3.0216 | -34.0929 | -5.4036
3.0277 | -34.0932 | -5.4152
3.0338 | -34.0936 | -5.4264
3.0399 | -34.0939 | -5.4377
3.046 | -34.0942 | -5.4491
3.0521 | -34.0945 | -5.4601
3.0582 | -34.0948 | -5.4719
3.0643 | -34.0951 | -5.4832
3.0704 | -34.0954 | -5.4948
3.0765 | -34.0957 | -5.5058
3.0827 | -34.0959 | -5.5169
3.0888 | -34.0962 | -5.528
3.0949 | -34.0966 | -5.5391
3.101 | -34.0969 | -5.5504
3.1071 | -34.0972 | -5.562
3.1132 | -34.0976 | -5.5736
3.1193 | -34.0979 | -5.5849
3.1254 | -34.0982 | -5.596
3.1315 | -34.0986 | -5.6077
3.1376 | -34.0988 | -5.6189
3.1437 | -34.0991 | -5.6303
3.1498 | -34.0994 | -5.6421
3.1559 | -34.0997 | -5.6534
3.162 | -34.1 | -5.6652
3.1681 | -34.1002 | -5.6769
3.1742 | -34.1005 | -5.6883
3.1803 | -34.1008 | -5.6998
3.1864 | -34.1011 | -5.7108
3.1925 | -34.1014 | -5.7222
3.1986 | -34.1016 | -5.7339
3.2047 | -34.1019 | -5.7457
3.2108 | -34.1021 | -5.7566
3.2169 | -34.1024 | -5.7681
3.223 | -34.1026 | -5.7796
3.2292 | -34.1028 | -5.7912
3.2353 | -34.103 | -5.8019
3.2414 | -34.1032 | -5.8132
3.2475 | -34.1034 | -5.8244
3.2536 | -34.1037 | -5.8356
3.2597 | -34.1039 | -5.8465
3.2658 | -34.1041 | -5.8573
3.2719 | -34.1043 | -5.8682
3.278 | -34.1046 | -5.8796
3.2841 | -34.1048 | -5.8906
3.2902 | -34.105 | -5.9018
3.2963 | -34.1052 | -5.9131
3.3024 | -34.1053 | -5.9241
3.3085 | -34.1056 | -5.9354
3.3146 | -34.1057 | -5.9465
3.3207 | -34.106 | -5.9578
3.3268 | -34.1062 | -5.9691
3.3329 | -34.1064 | -5.9797
3.339 | -34.1065 | -5.9904
3.3451 | -34.1067 | -6.0014
3.3512 | -34.1069 | -6.012
3.3573 | -34.1071 | -6.0231
3.3634 | -34.1072 | -6.0341
3.3696 | -34.1074 | -6.0453
3.3757 | -34.1076 | -6.056
3.3818 | -34.1078 | -6.0668
3.3879 | -34.1079 | -6.0778
3.394 | -34.1081 | -6.0891
3.4001 | -34.1083 | -6.0998
3.4062 | -34.1084 | -6.1107
3.4123 | -34.1085 | -6.1212
3.4184 | -34.1087 | -6.1323
3.4245 | -34.1089 | -6.143
3.4306 | -34.109 | -6.1537
3.4367 | -34.1092 | -6.1642
3.4428 | -34.1093 | -6.1745
3.4489 | -34.1095 | -6.1854
3.455 | -34.1096 | -6.1962
3.4611 | -34.1098 | -6.2073
3.4672 | -34.1099 | -6.2183
3.4733 | -34.1101 | -6.2293
3.4794 | -34.1103 | -6.2401
3.4855 | -34.1104 | -6.2505
3.4916 | -34.1106 | -6.2609
3.4977 | -34.1107 | -6.2721
3.5038 | -34.1109 | -6.2829
3.5099 | -34.111 | -6.2931
3.5161 | -34.1112 | -6.3035
3.5222 | -34.1113 | -6.314
3.5283 | -34.1114 | -6.3247
3.5344 | -34.1115 | -6.3355
3.5405 | -34.1117 | -6.346
3.5466 | -34.1119 | -6.3569
3.5527 | -34.112 | -6.3677
3.5588 | -34.1121 | -6.3783
3.5649 | -34.1121 | -6.3891
3.571 | -34.1123 | -6.3997
3.5771 | -34.1124 | -6.4107
3.5832 | -34.1125 | -6.4207
3.5893 | -34.1127 | -6.4313
3.5954 | -34.1129 | -6.4421
3.6015 | -34.1131 | -6.453
3.6076 | -34.1133 | -6.4635
3.6137 | -34.1134 | -6.4737
3.6198 | -34.1137 | -6.484
3.6259 | -34.1138 | -6.4949
3.632 | -34.114 | -6.5056
3.6381 | -34.1141 | -6.5163
3.6442 | -34.1143 | -6.5269
3.6503 | -34.1145 | -6.5372
3.6565 | -34.1147 | -6.5479
3.6626 | -34.1148 | -6.5585
3.6687 | -34.1151 | -6.5692
3.6748 | -34.1153 | -6.5796
3.6809 | -34.1155 | -6.5903
3.687 | -34.1157 | -6.6012
3.6931 | -34.1159 | -6.6115
3.6992 | -34.1161 | -6.6222
3.7053 | -34.1163 | -6.6327
3.7114 | -34.1165 | -6.6435
3.7175 | -34.1167 | -6.6541
3.7236 | -34.1169 | -6.6643
3.7297 | -34.1171 | -6.6751
3.7358 | -34.1173 | -6.6859
3.7419 | -34.1175 | -6.6968
3.748 | -34.1177 | -6.7073
3.7541 | -34.1179 | -6.718
3.7602 | -34.1182 | -6.7286
3.7663 | -34.1184 | -6.7391
3.7724 | -34.1186 | -6.7497
3.7785 | -34.1189 | -6.7602
3.7846 | -34.1191 | -6.7707
3.7907 | -34.1193 | -6.7816
3.7969 | -34.1196 | -6.7923
3.803 | -34.1198 | -6.8026
3.8091 | -34.12 | -6.8132
3.8152 | -34.1203 | -6.8238
3.8213 | -34.1205 | -6.8346
3.8274 | -34.1207 | -6.845
3.8335 | -34.121 | -6.8558
3.8396 | -34.1213 | -6.8664
3.8457 | -34.1216 | -6.8772
3.8518 | -34.1219 | -6.8877
3.8579 | -34.1222 | -6.8983
3.864 | -34.1225 | -6.9091
3.8701 | -34.1228 | -6.9201
3.8762 | -34.123 | -6.9306
3.8823 | -34.1232 | -6.9414
3.8884 | -34.1235 | -6.952
3.8945 | -34.1238 | -6.9629
3.9006 | -34.1241 | -6.9739
3.9067 | -34.1243 | -6.985
3.9128 | -34.1246 | -6.9957
3.9189 | -34.1249 | -7.0063
3.925 | -34.1251 | -7.0171
3.9311 | -34.1253 | -7.0273
3.9372 | -34.1256 | -7.0382
3.9434 | -34.1259 | -7.0489
3.9495 | -34.1262 | -7.0597
3.9556 | -34.1265 | -7.0702
3.9617 | -34.1268 | -7.0805
3.9678 | -34.1271 | -7.091
3.9739 | -34.1273 | -7.1016
3.98 | -34.1276 | -7.1124
3.9861 | -34.1278 | -7.123
3.9922 | -34.128 | -7.1338
3.9983 | -34.1283 | -7.1441
4.0044 | -34.1286 | -7.1544
4.0105 | -34.1288 | -7.165
4.0166 | -34.1291 | -7.1758
4.0227 | -34.1294 | -7.1864
4.0288 | -34.1296 | -7.1974
4.0349 | -34.1298 | -7.2079
4.041 | -34.1301 | -7.2185
4.0471 | -34.1303 | -7.2293
4.0532 | -34.1305 | -7.2396
4.0593 | -34.1308 | -7.2504
4.0654 | -34.1311 | -7.2608
4.0715 | -34.1313 | -7.2719
4.0776 | -34.1316 | -7.2825
4.0838 | -34.1319 | -7.2935
4.0899 | -34.1322 | -7.3039
4.096 | -34.1324 | -7.3145
4.1021 | -34.1326 | -7.325
4.1082 | -34.1329 | -7.3354
4.1143 | -34.1331 | -7.3457
4.1204 | -34.1334 | -7.356
4.1265 | -34.1336 | -7.3664
4.1326 | -34.1339 | -7.3768
4.1387 | -34.1342 | -7.3873
4.1448 | -34.1344 | -7.3979
4.1509 | -34.1346 | -7.4086
4.157 | -34.1348 | -7.4195
4.1631 | -34.1351 | -7.4302
4.1692 | -34.1353 | -7.4407
4.1753 | -34.1356 | -7.4512
4.1814 | -34.1357 | -7.4615
4.1875 | -34.136 | -7.4723
4.1936 | -34.1362 | -7.4829
4.1997 | -34.1365 | -7.4931
4.2058 | -34.1367 | -7.504
4.2119 | -34.1369 | -7.5144
4.218 | -34.1371 | -7.5246
4.2241 | -34.1373 | -7.535
4.2303 | -34.1376 | -7.5454
4.2364 | -34.1378 | -7.5559
4.2425 | -34.138 | -7.5664
4.2486 | -34.1381 | -7.5771
4.2547 | -34.1383 | -7.5874
4.2608 | -34.1385 | -7.598
4.2669 | -34.1387 | -7.6082
4.273 | -34.1389 | -7.6188
4.2791 | -34.1391 | -7.6293
4.2852 | -34.1393 | -7.64
4.2913 | -34.1395 | -7.6502
4.2974 | -34.1397 | -7.6609
4.3035 | -34.1399 | -7.6716
4.3096 | -34.1401 | -7.6823
4.3157 | -34.1403 | -7.6926
4.3218 | -34.1405 | -7.7029
4.3279 | -34.1407 | -7.7134
4.334 | -34.1409 | -7.7238
4.3401 | -34.141 | -7.7339
4.3462 | -34.1413 | -7.7444
4.3523 | -34.1415 | -7.7545
4.3584 | -34.1417 | -7.7646
4.3645 | -34.1419 | -7.7749
4.3707 | -34.1421 | -7.7851
4.3768 | -34.1423 | -7.7954
4.3829 | -34.1425 | -7.8061
4.389 | -34.1427 | -7.8164
4.3951 | -34.1429 | -7.8266
4.4012 | -34.143 | -7.8373
4.4073 | -34.1432 | -7.8475
4.4134 | -34.1434 | -7.8578
4.4195 | -34.1436 | -7.8681
4.4256 | -34.1439 | -7.8783
4.4317 | -34.1441 | -7.8886
4.4378 | -34.1443 | -7.8986
4.4439 | -34.1444 | -7.9086
4.45 | -34.1446 | -7.9189
4.4561 | -34.1447 | -7.9292
4.4622 | -34.1449 | -7.9395
4.4683 | -34.1451 | -7.9498
4.4744 | -34.1453 | -7.9605
4.4805 | -34.1455 | -7.9704
4.4866 | -34.1457 | -7.9807
4.4927 | -34.1459 | -7.9912
4.4988 | -34.1461 | -8.0014
4.5049 | -34.1463 | -8.0113
4.511 | -34.1465 | -8.0213
4.5172 | -34.1467 | -8.0316
4.5233 | -34.1469 | -8.0422
4.5294 | -34.1471 | -8.0523
4.5355 | -34.1473 | -8.0626
4.5416 | -34.1475 | -8.0728
4.5477 | -34.1478 | -8.083
4.5538 | -34.1479 | -8.0929
4.5599 | -34.1482 | -8.1031
4.566 | -34.1484 | -8.1134
4.5721 | -34.1486 | -8.1236
4.5782 | -34.1488 | -8.1341
4.5843 | -34.149 | -8.1444
4.5904 | -34.1493 | -8.1547
4.5965 | -34.1495 | -8.1648
4.6026 | -34.1497 | -8.1754
4.6087 | -34.1499 | -8.1856
4.6148 | -34.1501 | -8.1961
4.6209 | -34.1503 | -8.2063
4.627 | -34.1506 | -8.2167
4.6331 | -34.1507 | -8.2268
4.6392 | -34.1509 | -8.2371
4.6453 | -34.1512 | -8.2474
4.6514 | -34.1514 | -8.2577
4.6576 | -34.1516 | -8.2682
4.6637 | -34.1517 | -8.2782
4.6698 | -34.1519 | -8.2883
4.6759 | -34.1522 | -8.2989
4.682 | -34.1523 | -8.3091
4.6881 | -34.1525 | -8.3193
4.6942 | -34.1527 | -8.3295
4.7003 | -34.1529 | -8.3398
4.7064 | -34.1532 | -8.35
4.7125 | -34.1534 | -8.3605
4.7186 | -34.1536 | -8.3705
4.7247 | -34.1538 | -8.3809
4.7308 | -34.1541 | -8.3911
4.7369 | -34.1543 | -8.4015
4.743 | -34.1545 | -8.4121
4.7491 | -34.1547 | -8.4226
4.7552 | -34.1549 | -8.433
4.7613 | -34.1551 | -8.4434
4.7674 | -34.1553 | -8.4535
4.7735 | -34.1555 | -8.4642
4.7796 | -34.1558 | -8.4747
4.7857 | -34.156 | -8.4851
4.7918 | -34.1563 | -8.4952
4.7979 | -34.1565 | -8.5054
4.8041 | -34.1567 | -8.5157
4.8102 | -34.157 | -8.5258
4.8163 | -34.1572 | -8.536
4.8224 | -34.1574 | -8.5463
4.8285 | -34.1577 | -8.5563
4.8346 | -34.1579 | -8.5664
4.8407 | -34.1581 | -8.5768
4.8468 | -34.1584 | -8.5874
4.8529 | -34.1586 | -8.5979
4.859 | -34.1588 | -8.6085
4.8651 | -34.1591 | -8.6186
4.8712 | -34.1593 | -8.6288
4.8773 | -34.1595 | -8.6392
4.8834 | -34.1597 | -8.6495
4.8895 | -34.16 | -8.6596
4.8956 | -34.1602 | -8.6699
4.9017 | -34.1604 | -8.6801
4.9078 | -34.1606 | -8.6907
4.9139 | -34.1608 | -8.7007
4.92 | -34.1611 | -8.7108
4.9261 | -34.1614 | -8.7208
4.9322 | -34.1617 | -8.7311
4.9383 | -34.1619 | -8.7413
4.9445 | -34.1621 | -8.7516
4.9506 | -34.1624 | -8.7618
4.9567 | -34.1626 | -8.772
4.9628 | -34.1628 | -8.7822
4.9689 | -34.1631 | -8.7929
4.975 | -34.1633 | -8.8035
4.9811 | -34.1636 | -8.8135
4.9872 | -34.1638 | -8.8236
4.9933 | -34.1642 | -8.8342
4.9994 | -34.1644 | -8.8449
5.0055 | -34.1647 | -8.8552
5.0116 | -34.1649 | -8.8659
5.0177 | -34.1652 | -8.8764
5.0238 | -34.1655 | -8.887
5.0299 | -34.1658 | -8.8977
5.036 | -34.1661 | -8.908
5.0421 | -34.1663 | -8.9182
5.0482 | -34.1666 | -8.9286
5.0543 | -34.1669 | -8.9387
5.0604 | -34.1672 | -8.9491
5.0665 | -34.1675 | -8.9592
5.0726 | -34.1677 | -8.9698
5.0787 | -34.168 | -8.9803
5.0848 | -34.1683 | -8.9908
5.091 | -34.1685 | -9.001
5.0971 | -34.1688 | -9.011
5.1032 | -34.1691 | -9.0217
5.1093 | -34.1694 | -9.032
5.1154 | -34.1696 | -9.0422
5.1215 | -34.1699 | -9.0524
5.1276 | -34.1701 | -9.0625
5.1337 | -34.1705 | -9.0727
5.1398 | -34.1707 | -9.0828
5.1459 | -34.171 | -9.0934
5.152 | -34.1713 | -9.1039
5.1581 | -34.1716 | -9.1144
5.1642 | -34.1719 | -9.1247
5.1703 | -34.1721 | -9.135
5.1764 | -34.1724 | -9.1452
5.1825 | -34.1727 | -9.1554
5.1886 | -34.173 | -9.1654
5.1947 | -34.1732 | -9.1757
5.2008 | -34.1736 | -9.1862
5.2069 | -34.1738 | -9.1967
5.213 | -34.1741 | -9.2068
5.2191 | -34.1744 | -9.2172
5.2252 | -34.1747 | -9.2276
5.2314 | -34.1749 | -9.2382
5.2375 | -34.1752 | -9.2483
5.2436 | -34.1755 | -9.2586
5.2497 | -34.1758 | -9.2687
5.2558 | -34.1761 | -9.2794
5.2619 | -34.1764 | -9.29
5.268 | -34.1767 | -9.3005
5.2741 | -34.177 | -9.3109
5.2802 | -34.1773 | -9.3216
5.2863 | -34.1776 | -9.3318
5.2924 | -34.1779 | -9.3424
5.2985 | -34.1782 | -9.3528
5.3046 | -34.1785 | -9.3632
5.3107 | -34.1788 | -9.3737
5.3168 | -34.1791 | -9.3841
5.3229 | -34.1794 | -9.3945
5.329 | -34.1796 | -9.4049
5.3351 | -34.18 | -9.415
5.3412 | -34.1802 | -9.4253
5.3473 | -34.1805 | -9.4358
5.3534 | -34.1808 | -9.4462
5.3595 | -34.1811 | -9.4566
5.3656 | -34.1814 | -9.467
5.3717 | -34.1817 | -9.4776
5.3779 | -34.182 | -9.4883
5.384 | -34.1823 | -9.4987
5.3901 | -34.1826 | -9.5093
5.3962 | -34.1829 | -9.5195
5.4023 | -34.1832 | -9.5301
5.4084 | -34.1835 | -9.5404
5.4145 | -34.1838 | -9.5506
5.4206 | -34.1841 | -9.561
5.4267 | -34.1844 | -9.5717
5.4328 | -34.1847 | -9.5821
5.4389 | -34.185 | -9.5923
5.445 | -34.1853 | -9.603
5.4511 | -34.1856 | -9.6132
5.4572 | -34.1859 | -9.6237
5.4633 | -34.1862 | -9.6341
5.4694 | -34.1865 | -9.6445
5.4755 | -34.1868 | -9.6547
5.4816 | -34.1871 | -9.6652
5.4877 | -34.1874 | -9.6754
5.4938 | -34.1876 | -9.6856
5.4999 | -34.188 | -9.6959
5.506 | -34.1883 | -9.7062
5.5121 | -34.1886 | -9.7165
5.5183 | -34.1889 | -9.7266
5.5244 | -34.1891 | -9.7373
5.5305 | -34.1895 | -9.7475
5.5366 | -34.1898 | -9.7577
5.5427 | -34.19 | -9.7682
5.5488 | -34.1903 | -9.7786
5.5549 | -34.1906 | -9.789
5.561 | -34.1909 | -9.7996
5.5671 | -34.1912 | -9.8098
5.5732 | -34.1915 | -9.8204
5.5793 | -34.1918 | -9.8311
5.5854 | -34.1921 | -9.8414
5.5915 | -34.1924 | -9.8521
5.5976 | -34.1927 | -9.8625
5.6037 | -34.193 | -9.8729
5.6098 | -34.1933 | -9.8833
5.6159 | -34.1936 | -9.8935
5.622 | -34.1939 | -9.904
5.6281 | -34.1942 | -9.9144
5.6342 | -34.1945 | -9.9251
5.6403 | -34.1947 | -9.9355
5.6464 | -34.195 | -9.946
5.6525 | -34.1953 | -9.9565
5.6586 | -34.1956 | -9.9667
5.6648 | -34.196 | -9.9771
5.6709 | -34.1963 | -9.9877
5.677 | -34.1966 | -9.9981
5.6831 | -34.1969 | -10.0085
5.6892 | -34.1971 | -10.0191
5.6953 | -34.1974 | -10.0295
5.7014 | -34.1977 | -10.0399
5.7075 | -34.198 | -10.0503
5.7136 | -34.1983 | -10.0607
5.7197 | -34.1985 | -10.0711
5.7258 | -34.1988 | -10.0813
5.7319 | -34.1991 | -10.0917
5.738 | -34.1994 | -10.1021
5.7441 | -34.1997 | -10.1124
5.7502 | -34.2001 | -10.1226
5.7563 | -34.2003 | -10.133
5.7624 | -34.2007 | -10.1436
5.7685 | -34.201 | -10.154
5.7746 | -34.2012 | -10.1646
5.7807 | -34.2015 | -10.1751
5.7868 | -34.2017 | -10.1857
5.7929 | -34.202 | -10.196
5.799 | -34.2023 | -10.2062
5.8052 | -34.2026 | -10.2165
5.8113 | -34.2029 | -10.2268
5.8174 | -34.2032 | -10.237
5.8235 | -34.2035 | -10.2474
5.8296 | -34.2037 | -10.2579
5.8357 | -34.204 | -10.2682
5.8418 | -34.2043 | -10.2787
5.8479 | -34.2046 | -10.2892
5.854 | -34.2048 | -10.2997
5.8601 | -34.2052 | -10.31
5.8662 | -34.2054 | -10.3205
5.8723 | -34.2057 | -10.3307
5.8784 | -34.206 | -10.341
5.8845 | -34.2063 | -10.3513
5.8906 | -34.2065 | -10.3617
5.8967 | -34.2068 | -10.3719
5.9028 | -34.2071 | -10.3821
5.9089 | -34.2075 | -10.3923
5.915 | -34.2078 | -10.4025
5.9211 | -34.208 | -10.4128
5.9272 | -34.2083 | -10.423
5.9333 | -34.2086 | -10.4332
5.9394 | -34.2089 | -10.4435
5.9455 | -34.2092 | -10.4536
5.9517 | -34.2096 | -10.4638
5.9578 | -34.2099 | -10.4742
5.9639 | -34.2102 | -10.4843
5.97 | -34.2105 | -10.4947
5.9761 | -34.2108 | -10.5049
5.9822 | -34.2111 | -10.5153
5.9883 | -34.2114 | -10.5257
5.9944 | -34.2117 | -10.5361
6.0005 | -34.212 | -10.5465
6.0066 | -34.2123 | -10.5567
6.0127 | -34.2126 | -10.5667
6.0188 | -34.2129 | -10.5773
6.0249 | -34.2132 | -10.5876
6.031 | -34.2135 | -10.5977
6.0371 | -34.2138 | -10.6081
6.0432 | -34.2141 | -10.6183
6.0493 | -34.2144 | -10.6285
6.0554 | -34.2147 | -10.6386
6.0615 | -34.215 | -10.6489
6.0676 | -34.2153 | -10.6594
6.0737 | -34.2156 | -10.6698
6.0798 | -34.2159 | -10.6801
6.0859 | -34.2162 | -10.6903
6.0921 | -34.2165 | -10.7003
6.0982 | -34.2168 | -10.7106
6.1043 | -34.2171 | -10.7208
6.1104 | -34.2174 | -10.7313
6.1165 | -34.2177 | -10.7412
6.1226 | -34.218 | -10.7514
6.1287 | -34.2183 | -10.7616
6.1348 | -34.2187 | -10.7718
6.1409 | -34.219 | -10.7821
6.147 | -34.2192 | -10.7923
6.1531 | -34.2195 | -10.8025
6.1592 | -34.2198 | -10.8126
6.1653 | -34.2201 | -10.8227
6.1714 | -34.2204 | -10.8331
6.1775 | -34.2207 | -10.8434
6.1836 | -34.221 | -10.8536
6.1897 | -34.2213 | -10.8639
6.1958 | -34.2216 | -10.8743
6.2019 | -34.222 | -10.8846
6.208 | -34.2223 | -10.8949
6.2141 | -34.2226 | -10.9051
6.2202 | -34.2229 | -10.9154
6.2263 | -34.2232 | -10.9256
6.2325 | -34.2235 | -10.9356
6.2386 | -34.2238 | -10.9456
6.2447 | -34.2241 | -10.956
6.2508 | -34.2244 | -10.9662
6.2569 | -34.2248 | -10.9764
6.263 | -34.2251 | -10.9866
6.2691 | -34.2254 | -10.9967
6.2752 | -34.2256 | -11.007
6.2813 | -34.2259 | -11.0172
6.2874 | -34.2262 | -11.0273
6.2935 | -34.2265 | -11.0375
6.2996 | -34.2268 | -11.0476
6.3057 | -34.2271 | -11.0579
6.3118 | -34.2274 | -11.0682
6.3179 | -34.2277 | -11.0784
6.324 | -34.228 | -11.0886
6.3301 | -34.2283 | -11.0988
6.3362 | -34.2286 | -11.1089
6.3423 | -34.229 | -11.1193
6.3484 | -34.2293 | -11.1295
6.3545 | -34.2295 | -11.1395
6.3606 | -34.2298 | -11.1494
6.3667 | -34.2301 | -11.1595
6.3728 | -34.2304 | -11.1697
6.379 | -34.2307 | -11.1799
6.3851 | -34.231 | -11.1898
6.3912 | -34.2313 | -11.2
6.3973 | -34.2316 | -11.2101
6.4034 | -34.2319 | -11.2201
6.4095 | -34.2322 | -11.2303
6.4156 | -34.2325 | -11.2405
6.4217 | -34.2328 | -11.2506
6.4278 | -34.2331 | -11.2609
6.4339 | -34.2334 | -11.271
6.44 | -34.2337 | -11.281
6.4461 | -34.234 | -11.2914
6.4522 | -34.2343 | -11.3013
6.4583 | -34.2346 | -11.3116
6.4644 | -34.2349 | -11.3218
6.4705 | -34.2352 | -11.3317
6.4766 | -34.2355 | -11.3419
6.4827 | -34.2358 | -11.3521
6.4888 | -34.236 | -11.3621
6.4949 | -34.2363 | -11.372
6.501 | -34.2366 | -11.3823
6.5071 | -34.237 | -11.3922
6.5132 | -34.2373 | -11.4024
6.5194 | -34.2376 | -11.4126
6.5255 | -34.2379 | -11.4225
6.5316 | -34.2382 | -11.4328
6.5377 | -34.2384 | -11.443
6.5438 | -34.2387 | -11.453
6.5499 | -34.239 | -11.463
6.556 | -34.2393 | -11.473
6.5621 | -34.2396 | -11.4834
6.5682 | -34.2399 | -11.4935
6.5743 | -34.2403 | -11.5036
6.5804 | -34.2406 | -11.5138
6.5865 | -34.2409 | -11.524
6.5926 | -34.2412 | -11.534
6.5987 | -34.2415 | -11.5439
6.6048 | -34.2418 | -11.554
6.6109 | -34.2421 | -11.5642
6.617 | -34.2424 | -11.5744
6.6231 | -34.2427 | -11.5845
6.6292 | -34.243 | -11.5947
6.6353 | -34.2433 | -11.605
6.6414 | -34.2436 | -11.6151
6.6475 | -34.2439 | -11.6252
6.6536 | -34.2442 | -11.6353
6.6597 | -34.2445 | -11.6452
6.6659 | -34.2448 | -11.6553
6.672 | -34.2451 | -11.6653
6.6781 | -34.2454 | -11.6754
6.6842 | -34.2457 | -11.6854
6.6903 | -34.246 | -11.6954
6.6964 | -34.2464 | -11.7055
6.7025 | -34.2466 | -11.7154
6.7086 | -34.2469 | -11.7256
6.7147 | -34.2472 | -11.7357
6.7208 | -34.2475 | -11.7459
6.7269 | -34.2478 | -11.7559
6.733 | -34.2481 | -11.7658
6.7391 | -34.2484 | -11.7759
6.7452 | -34.2487 | -11.7858
6.7513 | -34.249 | -11.7957
6.7574 | -34.2494 | -11.8058
6.7635 | -34.2496 | -11.8157
6.7696 | -34.25 | -11.8257
6.7757 | -34.2503 | -11.8357
6.7818 | -34.2506 | -11.8458
6.7879 | -34.2509 | -11.8557
6.794 | -34.2512 | -11.8656
6.8001 | -34.2515 | -11.8757
6.8063 | -34.2518 | -11.8857
6.8124 | -34.2521 | -11.896
6.8185 | -34.2524 | -11.9062
6.8246 | -34.2527 | -11.9161
6.8307 | -34.253 | -11.9261
6.8368 | -34.2533 | -11.9363
6.8429 | -34.2536 | -11.9464
6.849 | -34.2539 | -11.9565
6.8551 | -34.2543 | -11.9664
6.8612 | -34.2546 | -11.9765
6.8673 | -34.2549 | -11.9865
6.8734 | -34.2552 | -11.9967
6.8795 | -34.2555 | -12.0067
6.8856 | -34.2558 | -12.0168
6.8917 | -34.2561 | -12.0268
6.8978 | -34.2564 | -12.0366
6.9039 | -34.2567 | -12.0464
6.91 | -34.257 | -12.0565
6.9161 | -34.2574 | -12.0664
6.9222 | -34.2576 | -12.0764
6.9283 | -34.258 | -12.0861
6.9344 | -34.2583 | -12.0962
6.9405 | -34.2586 | -12.1061
6.9466 | -34.259 | -12.1163
6.9528 | -34.2593 | -12.1264
6.9589 | -34.2596 | -12.1365
6.965 | -34.26 | -12.1466
6.9711 | -34.2603 | -12.1568
6.9772 | -34.2606 | -12.1667
6.9833 | -34.2609 | -12.1767
6.9894 | -34.2613 | -12.1869
6.9955 | -34.2616 | -12.1967
7.0016 | -34.2619 | -12.2067
7.0077 | -34.2622 | -12.2166
7.0138 | -34.2625 | -12.2266
7.0199 | -34.2628 | -12.2365
7.026 | -34.2632 | -12.2464
7.0321 | -34.2635 | -12.2562
7.0382 | -34.2638 | -12.266
7.0443 | -34.2642 | -12.2761
7.0504 | -34.2645 | -12.2861
7.0565 | -34.2648 | -12.2961
7.0626 | -34.2651 | -12.3059
7.0687 | -34.2654 | -12.3155
7.0748 | -34.2658 | -12.3252
7.0809 | -34.2661 | -12.3351
7.087 | -34.2664 | -12.3452
7.0932 | -34.2667 | -12.3552
7.0993 | -34.267 | -12.3652
7.1054 | -34.2674 | -12.375
7.1115 | -34.2677 | -12.3848
7.1176 | -34.268 | -12.3947
7.1237 | -34.2683 | -12.4046
7.1298 | -34.2686 | -12.4145
7.1359 | -34.2689 | -12.4244
7.142 | -34.2693 | -12.4341
7.1481 | -34.2696 | -12.4439
7.1542 | -34.2699 | -12.4538
7.1603 | -34.2702 | -12.4637
7.1664 | -34.2706 | -12.4736
7.1725 | -34.2709 | -12.4835
7.1786 | -34.2713 | -12.4934
7.1847 | -34.2715 | -12.5034
7.1908 | -34.2719 | -12.5135
7.1969 | -34.2722 | -12.5235
7.203 | -34.2725 | -12.5332
7.2091 | -34.2728 | -12.543
7.2152 | -34.2732 | -12.5529
7.2213 | -34.2735 | -12.5627
7.2274 | -34.2738 | -12.5726
7.2335 | -34.2741 | -12.5824
7.2397 | -34.2745 | -12.5922
7.2458 | -34.2748 | -12.6019
7.2519 | -34.2752 | -12.6118
7.258 | -34.2755 | -12.6218
7.2641 | -34.2758 | -12.6317
7.2702 | -34.2761 | -12.6418
7.2763 | -34.2765 | -12.6517
7.2824 | -34.2768 | -12.6616
7.2885 | -34.2772 | -12.6715
7.2946 | -34.2775 | -12.6814
7.3007 | -34.2778 | -12.6914
7.3068 | -34.2781 | -12.7014
7.3129 | -34.2785 | -12.7114
7.319 | -34.2788 | -12.7212
7.3251 | -34.2791 | -12.731
7.3312 | -34.2794 | -12.7408
7.3373 | -34.2798 | -12.7506
7.3434 | -34.2801 | -12.7603
7.3495 | -34.2805 | -12.77
7.3556 | -34.2808 | -12.78
7.3617 | -34.2812 | -12.79
7.3678 | -34.2815 | -12.7998
7.3739 | -34.2818 | -12.8098
7.3801 | -34.2822 | -12.8199
7.3862 | -34.2825 | -12.8297
7.3923 | -34.2829 | -12.8397
7.3984 | -34.2832 | -12.8494
7.4045 | -34.2835 | -12.8592
7.4106 | -34.2839 | -12.869
7.4167 | -34.2842 | -12.879
7.4228 | -34.2846 | -12.8889
7.4289 | -34.2849 | -12.8988
7.435 | -34.2852 | -12.9084
7.4411 | -34.2856 | -12.9183
7.4472 | -34.2859 | -12.9282
7.4533 | -34.2863 | -12.9378
7.4594 | -34.2866 | -12.9479
7.4655 | -34.2869 | -12.9576
7.4716 | -34.2873 | -12.9676
7.4777 | -34.2876 | -12.9775
7.4838 | -34.2879 | -12.9871
7.4899 | -34.2883 | -12.9969
7.496 | -34.2886 | -13.0065
7.5021 | -34.289 | -13.0163
7.5082 | -34.2893 | -13.0262
7.5143 | -34.2896 | -13.0361
7.5204 | -34.29 | -13.0461
7.5266 | -34.2903 | -13.0557
7.5327 | -34.2906 | -13.0655
7.5388 | -34.291 | -13.0755
7.5449 | -34.2913 | -13.0856
7.551 | -34.2917 | -13.0953
7.5571 | -34.292 | -13.1052
7.5632 | -34.2924 | -13.1151
7.5693 | -34.2927 | -13.1248
7.5754 | -34.2931 | -13.1345
7.5815 | -34.2934 | -13.1444
7.5876 | -34.2937 | -13.1544
7.5937 | -34.2941 | -13.1643
7.5998 | -34.2945 | -13.1741
7.6059 | -34.2948 | -13.1841
7.612 | -34.2952 | -13.194
7.6181 | -34.2955 | -13.2037
7.6242 | -34.2959 | -13.2138
7.6303 | -34.2962 | -13.2237
7.6364 | -34.2965 | -13.2337
7.6425 | -34.2969 | -13.2436
7.6486 | -34.2972 | -13.2535
7.6547 | -34.2976 | -13.2632
7.6608 | -34.2979 | -13.273
7.667 | -34.2982 | -13.2829
7.6731 | -34.2986 | -13.2928
7.6792 | -34.2989 | -13.3029
7.6853 | -34.2993 | -13.3128
7.6914 | -34.2996 | -13.3228
7.6975 | -34.3 | -13.3327
7.7036 | -34.3003 | -13.3425
7.7097 | -34.3007 | -13.3525
7.7158 | -34.3011 | -13.3624
7.7219 | -34.3014 | -13.3721
7.728 | -34.3017 | -13.3821
7.7341 | -34.3021 | -13.392
7.7402 | -34.3025 | -13.4019
7.7463 | -34.3028 | -13.4118
7.7524 | -34.3031 | -13.4219
7.7585 | -34.3035 | -13.4318
7.7646 | -34.3038 | -13.4417
7.7707 | -34.3042 | -13.4516
7.7768 | -34.3045 | -13.4617
7.7829 | -34.3048 | -13.4718
7.789 | -34.3052 | -13.4819
7.7951 | -34.3055 | -13.4918
7.8012 | -34.3059 | -13.5016
7.8073 | -34.3063 | -13.5117
7.8135 | -34.3066 | -13.5215
7.8196 | -34.307 | -13.5316
7.8257 | -34.3073 | -13.5416
7.8318 | -34.3076 | -13.5516
7.8379 | -34.308 | -13.5614
7.844 | -34.3084 | -13.5713
7.8501 | -34.3087 | -13.5813
7.8562 | -34.309 | -13.5914
7.8623 | -34.3094 | -13.6012
7.8684 | -34.3098 | -13.6112
7.8745 | -34.3101 | -13.6212
7.8806 | -34.3104 | -13.6312
7.8867 | -34.3108 | -13.641
7.8928 | -34.3112 | -13.6506
7.8989 | -34.3115 | -13.6606
7.905 | -34.3119 | -13.6705
7.9111 | -34.3122 | -13.6805
7.9172 | -34.3126 | -13.6904
7.9233 | -34.3129 | -13.7005
7.9294 | -34.3132 | -13.7105
7.9355 | -34.3136 | -13.7202
7.9416 | -34.3139 | -13.7302
7.9477 | -34.3142 | -13.7399
7.9539 | -34.3146 | -13.7497
7.96 | -34.3149 | -13.7597
7.9661 | -34.3153 | -13.7696
7.9722 | -34.3157 | -13.7794
7.9783 | -34.316 | -13.7895
7.9844 | -34.3164 | -13.7993
7.9905 | -34.3167 | -13.8094
7.9966 | -34.3171 | -13.8194
8.0027 | -34.3174 | -13.8293
8.0088 | -34.3177 | -13.8392
8.0149 | -34.3181 | -13.8492
8.021 | -34.3184 | -13.8592
8.0271 | -34.3188 | -13.8693
8.0332 | -34.3191 | -13.8791
8.0393 | -34.3195 | -13.8892
8.0454 | -34.3198 | -13.8993
8.0515 | -34.3202 | -13.9093
8.0576 | -34.3206 | -13.9194
8.0637 | -34.3209 | -13.9292
8.0698 | -34.3213 | -13.939
8.0759 | -34.3216 | -13.9489
8.082 | -34.3219 | -13.9588
8.0881 | -34.3223 | -13.9687
8.0942 | -34.3227 | -13.9787
8.1004 | -34.3231 | -13.9887
8.1065 | -34.3234 | -13.9985
8.1126 | -34.3237 | -14.0086
8.1187 | -34.3241 | -14.0186
8.1248 | -34.3244 | -14.0283
8.1309 | -34.3248 | -14.0381
8.137 | -34.3251 | -14.0479
8.1431 | -34.3255 | -14.058
8.1492 | -34.3258 | -14.0679
8.1553 | -34.3262 | -14.0779
8.1614 | -34.3265 | -14.088
8.1675 | -34.3269 | -14.0979
8.1736 | -34.3272 | -14.1079
8.1797 | -34.3276 | -14.1179
8.1858 | -34.3279 | -14.1279
8.1919 | -34.3283 | -14.1378
8.198 | -34.3286 | -14.1476
8.2041 | -34.329 | -14.1576
8.2102 | -34.3293 | -14.1675
8.2163 | -34.3297 | -14.1773
8.2224 | -34.33 | -14.1872
8.2285 | -34.3304 | -14.1969
8.2346 | -34.3307 | -14.2068
8.2408 | -34.3311 | -14.2167
8.2469 | -34.3314 | -14.2267
8.253 | -34.3318 | -14.2365
8.2591 | -34.3322 | -14.2465
8.2652 | -34.3325 | -14.2564
8.2713 | -34.3329 | -14.2664
8.2774 | -34.3333 | -14.2764
8.2835 | -34.3336 | -14.2865
8.2896 | -34.334 | -14.2963
8.2957 | -34.3343 | -14.3063
8.3018 | -34.3347 | -14.3162
8.3079 | -34.335 | -14.3261
8.314 | -34.3354 | -14.336
8.3201 | -34.3358 | -14.3458
8.3262 | -34.3361 | -14.3558
8.3323 | -34.3365 | -14.3657
8.3384 | -34.3368 | -14.3757
8.3445 | -34.3372 | -14.3855
8.3506 | -34.3375 | -14.3955
8.3567 | -34.3379 | -14.4054
8.3628 | -34.3383 | -14.4154
8.3689 | -34.3386 | -14.4252
8.375 | -34.339 | -14.4351
8.3812 | -34.3393 | -14.4449
8.3873 | -34.3397 | -14.4548
8.3934 | -34.34 | -14.4648
8.3995 | -34.3404 | -14.4746
8.4056 | -34.3407 | -14.4845
8.4117 | -34.3411 | -14.4944
8.4178 | -34.3414 | -14.5043
8.4239 | -34.3418 | -14.5141
8.43 | -34.3421 | -14.5241
8.4361 | -34.3424 | -14.534
8.4422 | -34.3428 | -14.5439
8.4483 | -34.3432 | -14.5537
8.4544 | -34.3436 | -14.5636
8.4605 | -34.3439 | -14.5734
8.4666 | -34.3442 | -14.5834
8.4727 | -34.3446 | -14.5933
8.4788 | -34.3449 | -14.6032
8.4849 | -34.3453 | -14.6131
8.491 | -34.3456 | -14.623
8.4971 | -34.346 | -14.6327
8.5032 | -34.3464 | -14.6425
8.5093 | -34.3467 | -14.6522
8.5154 | -34.3471 | -14.6622
8.5215 | -34.3475 | -14.6719
8.5277 | -34.3478 | -14.6819
8.5338 | -34.3482 | -14.6917
8.5399 | -34.3485 | -14.7015
8.546 | -34.3489 | -14.7112
8.5521 | -34.3493 | -14.7211
8.5582 | -34.3496 | -14.731
8.5643 | -34.35 | -14.7406
8.5704 | -34.3503 | -14.7502
8.5765 | -34.3507 | -14.7602
8.5826 | -34.351 | -14.7698
8.5887 | -34.3514 | -14.7796
8.5948 | -34.3518 | -14.7894
8.6009 | -34.3522 | -14.7991
8.607 | -34.3525 | -14.809
8.6131 | -34.3529 | -14.8189
8.6192 | -34.3533 | -14.8286
8.6253 | -34.3536 | -14.8384
8.6314 | -34.354 | -14.8481
8.6375 | -34.3543 | -14.858
8.6436 | -34.3547 | -14.8678
8.6497 | -34.3551 | -14.8776
8.6558 | -34.3554 | -14.8874
8.6619 | -34.3558 | -14.8972
8.6681 | -34.3561 | -14.9072
8.6742 | -34.3565 | -14.9169
8.6803 | -34.3569 | -14.9267
8.6864 | -34.3572 | -14.9363
8.6925 | -34.3576 | -14.9459
8.6986 | -34.358 | -14.9558
8.7047 | -34.3583 | -14.9657
8.7108 | -34.3587 | -14.9755
8.7169 | -34.359 | -14.9853
8.723 | -34.3594 | -14.995
8.7291 | -34.3598 | -15.0047
8.7352 | -34.3601 | -15.0146
8.7413 | -34.3605 | -15.0245
8.7474 | -34.3608 | -15.0342
8.7535 | -34.3612 | -15.0439
8.7596 | -34.3616 | -15.0538
8.7657 | -34.362 | -15.0637
8.7718 | -34.3623 | -15.0736
8.7779 | -34.3627 | -15.0834
8.784 | -34.363 | -15.0933
8.7901 | -34.3635 | -15.1031
8.7962 | -34.3638 | -15.1127
8.8023 | -34.3642 | -15.1224
8.8084 | -34.3645 | -15.1321
8.8146 | -34.3649 | -15.1419
8.8207 | -34.3652 | -15.1517
8.8268 | -34.3656 | -15.1612
8.8329 | -34.366 | -15.171
8.839 | -34.3664 | -15.1807
8.8451 | -34.3667 | -15.1904
8.8512 | -34.3671 | -15.2001
8.8573 | -34.3675 | -15.2099
8.8634 | -34.3678 | -15.2194
8.8695 | -34.3682 | -15.2294
8.8756 | -34.3686 | -15.2391
8.8817 | -34.3689 | -15.2488
8.8878 | -34.3693 | -15.2587
8.8939 | -34.3697 | -15.2685
8.9 | -34.3701 | -15.2782
8.9061 | -34.3704 | -15.2879
8.9122 | -34.3708 | -15.2976
8.9183 | -34.3712 | -15.3074
8.9244 | -34.3716 | -15.3173
8.9305 | -34.372 | -15.3269
8.9366 | -34.3724 | -15.3366
8.9427 | -34.3727 | -15.3464
8.9488 | -34.3731 | -15.3561
8.955 | -34.3735 | -15.3658
8.9611 | -34.3738 | -15.3757
8.9672 | -34.3743 | -15.3853
8.9733 | -34.3746 | -15.3949
8.9794 | -34.375 | -15.4045
8.9855 | -34.3754 | -15.4144
8.9916 | -34.3758 | -15.4242
8.9977 | -34.3761 | -15.4339
9.0038 | -34.3766 | -15.4435
9.0099 | -34.3769 | -15.4532
9.016 | -34.3773 | -15.463
9.0221 | -34.3777 | -15.4725
9.0282 | -34.3781 | -15.4821
9.0343 | -34.3784 | -15.4918
9.0404 | -34.3789 | -15.5015
9.0465 | -34.3792 | -15.5113
9.0526 | -34.3796 | -15.521
9.0587 | -34.38 | -15.5306
9.0648 | -34.3804 | -15.5402
9.0709 | -34.3808 | -15.5499
9.077 | -34.3812 | -15.5597
9.0831 | -34.3816 | -15.5695
9.0892 | -34.3819 | -15.5792
9.0953 | -34.3824 | -15.589
9.1015 | -34.3827 | -15.5987
9.1076 | -34.3831 | -15.6083
9.1137 | -34.3835 | -15.6179
9.1198 | -34.3839 | -15.6276
9.1259 | -34.3843 | -15.6372
9.132 | -34.3847 | -15.6471
9.1381 | -34.3851 | -15.6567
9.1442 | -34.3855 | -15.6661
9.1503 | -34.3859 | -15.6758
9.1564 | -34.3863 | -15.6854
9.1625 | -34.3866 | -15.6952
9.1686 | -34.387 | -15.7048
9.1747 | -34.3874 | -15.7145
9.1808 | -34.3879 | -15.7241
9.1869 | -34.3882 | -15.7338
9.193 | -34.3886 | -15.7435
9.1991 | -34.389 | -15.7532
9.2052 | -34.3894 | -15.7629
9.2113 | -34.3898 | -15.7726
9.2174 | -34.3902 | -15.7823
9.2235 | -34.3906 | -15.7919
9.2296 | -34.391 | -15.8016
9.2357 | -34.3914 | -15.8114
9.2419 | -34.3918 | -15.821
9.248 | -34.3923 | -15.8306
9.2541 | -34.3926 | -15.8403
9.2602 | -34.393 | -15.8501
9.2663 | -34.3934 | -15.8595
9.2724 | -34.3938 | -15.8691
9.2785 | -34.3942 | -15.8787
9.2846 | -34.3946 | -15.8883
9.2907 | -34.395 | -15.8979
9.2968 | -34.3954 | -15.9077
9.3029 | -34.3958 | -15.9174
9.309 | -34.3962 | -15.927
9.3151 | -34.3966 | -15.9365
9.3212 | -34.397 | -15.946
9.3273 | -34.3974 | -15.9557
9.3334 | -34.3978 | -15.9652
9.3395 | -34.3982 | -15.9747
9.3456 | -34.3986 | -15.9842
9.3517 | -34.399 | -15.994
9.3578 | -34.3994 | -16.0036
9.3639 | -34.3998 | -16.0131
9.37 | -34.4002 | -16.0227
9.3761 | -34.4006 | -16.0325
9.3822 | -34.401 | -16.0422
9.3884 | -34.4015 | -16.0516
9.3945 | -34.4019 | -16.0613
9.4006 | -34.4023 | -16.0709
9.4067 | -34.4027 | -16.0805
9.4128 | -34.4031 | -16.0902
9.4189 | -34.4035 | -16.1
9.425 | -34.4039 | -16.1095
9.4311 | -34.4043 | -16.119
9.4372 | -34.4047 | -16.1286
9.4433 | -34.4051 | -16.1382
9.4494 | -34.4055 | -16.1477
9.4555 | -34.406 | -16.1573
9.4616 | -34.4064 | -16.1669
9.4677 | -34.4068 | -16.1765
9.4738 | -34.4072 | -16.1861
9.4799 | -34.4076 | -16.1957
9.486 | -34.408 | -16.2053
9.4921 | -34.4084 | -16.2149
9.4982 | -34.4088 | -16.2242
9.5043 | -34.4092 | -16.2338
9.5104 | -34.4096 | -16.2433
9.5165 | -34.41 | -16.2529
9.5226 | -34.4104 | -16.2626
9.5288 | -34.4108 | -16.272
9.5349 | -34.4112 | -16.2817
9.541 | -34.4117 | -16.291
9.5471 | -34.412 | -16.3006
9.5532 | -34.4125 | -16.3103
9.5593 | -34.4129 | -16.3199
9.5654 | -34.4133 | -16.3295
9.5715 | -34.4137 | -16.339
9.5776 | -34.4141 | -16.3487
9.5837 | -34.4145 | -16.3582
9.5898 | -34.4148 | -16.3679
9.5959 | -34.4153 | -16.3775
9.602 | -34.4157 | -16.3869
9.6081 | -34.4161 | -16.3964
9.6142 | -34.4165 | -16.4059
9.6203 | -34.4169 | -16.4155
9.6264 | -34.4173 | -16.425
9.6325 | -34.4177 | -16.4346
9.6386 | -34.4181 | -16.4441
9.6447 | -34.4185 | -16.4537
9.6508 | -34.4189 | -16.4632
9.6569 | -34.4193 | -16.4727
9.663 | -34.4197 | -16.4823
9.6691 | -34.4201 | -16.4919
9.6753 | -34.4206 | -16.5014
9.6814 | -34.421 | -16.5108
9.6875 | -34.4214 | -16.5201
9.6936 | -34.4218 | -16.5296
9.6997 | -34.4222 | -16.5392
9.7058 | -34.4226 | -16.5486
9.7119 | -34.423 | -16.5583
9.718 | -34.4234 | -16.5679
9.7241 | -34.4239 | -16.5774
9.7302 | -34.4242 | -16.5869
9.7363 | -34.4247 | -16.5963
9.7424 | -34.4251 | -16.6057
9.7485 | -34.4255 | -16.6153
9.7546 | -34.4259 | -16.6246
9.7607 | -34.4263 | -16.6344
9.7668 | -34.4267 | -16.644
9.7729 | -34.4271 | -16.6534
9.779 | -34.4275 | -16.6628
9.7851 | -34.428 | -16.6724
9.7912 | -34.4284 | -16.6819
9.7973 | -34.4288 | -16.6915
9.8034 | -34.4292 | -16.7008
9.8095 | -34.4296 | -16.7105
9.8157 | -34.4301 | -16.72
9.8218 | -34.4305 | -16.7296
9.8279 | -34.4309 | -16.739
9.834 | -34.4313 | -16.7486
9.8401 | -34.4317 | -16.7581
9.8462 | -34.4321 | -16.7673
9.8523 | -34.4325 | -16.7768
9.8584 | -34.4329 | -16.7863
9.8645 | -34.4333 | -16.796
9.8706 | -34.4337 | -16.8056
9.8767 | -34.4341 | -16.815
9.8828 | -34.4345 | -16.8244
9.8889 | -34.4349 | -16.8337
9.895 | -34.4353 | -16.8433
9.9011 | -34.4357 | -16.8528
9.9072 | -34.4361 | -16.8622
9.9133 | -34.4365 | -16.8717
9.9194 | -34.4369 | -16.8811
9.9255 | -34.4373 | -16.8906
9.9316 | -34.4377 | -16.9
9.9377 | -34.4381 | -16.9095
9.9438 | -34.4385 | -16.919
9.9499 | -34.4389 | -16.9287
9.956 | -34.4393 | -16.9381
9.9622 | -34.4397 | -16.9477
9.9683 | -34.4401 | -16.9572
9.9744 | -34.4405 | -16.9669
9.9805 | -34.441 | -16.9763
9.9866 | -34.4414 | -16.9857
9.9927 | -34.4418 | -16.9951
9.9988 | -34.4422 | -17.0047
10.0049 | -34.4426 | -17.0141
10.011 | -34.443 | -17.0236
10.0171 | -34.4434 | -17.033
10.0232 | -34.4438 | -17.0424
10.0293 | -34.4442 | -17.0519
10.0354 | -34.4447 | -17.0613
10.0415 | -34.4451 | -17.0708
10.0476 | -34.4455 | -17.0804
10.0537 | -34.4459 | -17.09
10.0598 | -34.4463 | -17.0994
10.0659 | -34.4467 | -17.1089
10.072 | -34.4471 | -17.1182
10.0781 | -34.4475 | -17.1276
10.0842 | -34.4479 | -17.137
10.0903 | -34.4483 | -17.1465
10.0964 | -34.4487 | -17.1556
10.1026 | -34.4491 | -17.1651
10.1087 | -34.4495 | -17.1745
10.1148 | -34.4499 | -17.184
10.1209 | -34.4503 | -17.1934
10.127 | -34.4507 | -17.2027
10.1331 | -34.4511 | -17.212
10.1392 | -34.4515 | -17.2214
10.1453 | -34.4519 | -17.2308
10.1514 | -34.4524 | -17.2402
10.1575 | -34.4527 | -17.2497
10.1636 | -34.4531 | -17.2589
10.1697 | -34.4536 | -17.2684
10.1758 | -34.454 | -17.2777
10.1819 | -34.4544 | -17.2871
10.188 | -34.4548 | -17.2966
10.1941 | -34.4552 | -17.306
10.2002 | -34.4556 | -17.3153
10.2063 | -34.456 | -17.3247
10.2124 | -34.4564 | -17.3341
10.2185 | -34.4568 | -17.3436
10.2246 | -34.4572 | -17.3529
10.2307 | -34.4576 | -17.3623
10.2368 | -34.458 | -17.3715
10.2429 | -34.4584 | -17.3808
10.2491 | -34.4588 | -17.3903
10.2552 | -34.4592 | -17.3997
10.2613 | -34.4596 | -17.4091
10.2674 | -34.46 | -17.4186
10.2735 | -34.4604 | -17.4282
10.2796 | -34.4608 | -17.4375
10.2857 | -34.4612 | -17.447
10.2918 | -34.4616 | -17.4565
10.2979 | -34.462 | -17.4659
10.304 | -34.4625 | -17.4752
10.3101 | -34.4629 | -17.4846
10.3162 | -34.4633 | -17.4941
10.3223 | -34.4637 | -17.5034
10.3284 | -34.4641 | -17.5127
10.3345 | -34.4645 | -17.522
10.3406 | -34.4649 | -17.5313
10.3467 | -34.4653 | -17.5406
10.3528 | -34.4657 | -17.5499
10.3589 | -34.4661 | -17.5593
10.365 | -34.4665 | -17.5688
10.3711 | -34.4669 | -17.5783
10.3772 | -34.4673 | -17.5876
10.3833 | -34.4677 | -17.597
10.3895 | -34.4681 | -17.6064
10.3956 | -34.4685 | -17.6157
10.4017 | -34.4689 | -17.6249
10.4078 | -34.4693 | -17.6343
10.4139 | -34.4697 | -17.6436
10.42 | -34.4701 | -17.6531
10.4261 | -34.4705 | -17.6624
10.4322 | -34.4709 | -17.6717
10.4383 | -34.4713 | -17.6811
10.4444 | -34.4717 | -17.6904
10.4505 | -34.4721 | -17.6998
10.4566 | -34.4725 | -17.7091
10.4627 | -34.4729 | -17.7183
10.4688 | -34.4733 | -17.7275
10.4749 | -34.4737 | -17.7369
10.481 | -34.4741 | -17.7462
10.4871 | -34.4746 | -17.7555
10.4932 | -34.475 | -17.7647
10.4993 | -34.4754 | -17.7741
10.5054 | -34.4758 | -17.7835
10.5115 | -34.4762 | -17.7929
10.5176 | -34.4766 | -17.8021
10.5237 | -34.477 | -17.8115
10.5298 | -34.4774 | -17.8208
10.536 | -34.4778 | -17.8302
10.5421 | -34.4782 | -17.8393
10.5482 | -34.4786 | -17.8487
10.5543 | -34.4791 | -17.858
10.5604 | -34.4795 | -17.8671
10.5665 | -34.4799 | -17.8763
10.5726 | -34.4803 | -17.8855
10.5787 | -34.4807 | -17.8948
10.5848 | -34.4811 | -17.9039
10.5909 | -34.4815 | -17.9133
10.597 | -34.4819 | -17.9227
10.6031 | -34.4823 | -17.9322
10.6092 | -34.4827 | -17.9415
10.6153 | -34.4831 | -17.9509
10.6214 | -34.4835 | -17.96
10.6275 | -34.4839 | -17.9692
10.6336 | -34.4843 | -17.9785
10.6397 | -34.4847 | -17.9879
10.6458 | -34.4851 | -17.9971
10.6519 | -34.4855 | -18.0065
10.658 | -34.4859 | -18.0157
10.6641 | -34.4863 | -18.0251
10.6702 | -34.4867 | -18.0344
10.6764 | -34.4871 | -18.0437
10.6825 | -34.4875 | -18.0529
10.6886 | -34.4879 | -18.0621
10.6947 | -34.4883 | -18.0714
10.7008 | -34.4888 | -18.0807
10.7069 | -34.4892 | -18.0901
10.713 | -34.4896 | -18.0994
10.7191 | -34.49 | -18.1086
10.7252 | -34.4904 | -18.1177
10.7313 | -34.4908 | -18.1269
10.7374 | -34.4912 | -18.1363
10.7435 | -34.4916 | -18.1456
10.7496 | -34.492 | -18.1547
10.7557 | -34.4924 | -18.164
10.7618 | -34.4928 | -18.173
10.7679 | -34.4932 | -18.1823
10.774 | -34.4936 | -18.1916
10.7801 | -34.494 | -18.2009
10.7862 | -34.4944 | -18.21
10.7923 | -34.4948 | -18.2193
10.7984 | -34.4953 | -18.2285
10.8045 | -34.4957 | -18.2377
10.8106 | -34.4961 | -18.247
10.8168 | -34.4965 | -18.2561
10.8229 | -34.4969 | -18.2654
10.829 | -34.4973 | -18.2746
10.8351 | -34.4977 | -18.2836
10.8412 | -34.4982 | -18.2927
10.8473 | -34.4986 | -18.3018
10.8534 | -34.499 | -18.3111
10.8595 | -34.4994 | -18.3204
10.8656 | -34.4998 | -18.3294
10.8717 | -34.5002 | -18.3386
10.8778 | -34.5007 | -18.3479
10.8839 | -34.5011 | -18.3571
10.89 | -34.5015 | -18.3663
10.8961 | -34.5019 | -18.3755
10.9022 | -34.5023 | -18.3846
10.9083 | -34.5027 | -18.3938
10.9144 | -34.5031 | -18.4028
10.9205 | -34.5035 | -18.4121
10.9266 | -34.5039 | -18.4212
10.9327 | -34.5044 | -18.4303
10.9388 | -34.5048 | -18.4394
10.9449 | -34.5052 | -18.4486
10.951 | -34.5056 | -18.4578
10.9571 | -34.506 | -18.467
10.9633 | -34.5064 | -18.4763
10.9694 | -34.5068 | -18.4856
10.9755 | -34.5073 | -18.4948
10.9816 | -34.5077 | -18.5041
10.9877 | -34.5081 | -18.5131
10.9938 | -34.5085 | -18.5223
10.9999 | -34.5089 | -18.5314
11.006 | -34.5093 | -18.5403
11.0121 | -34.5098 | -18.5495
11.0182 | -34.5102 | -18.5586
11.0243 | -34.5106 | -18.5678
11.0304 | -34.511 | -18.577
11.0365 | -34.5114 | -18.5861
11.0426 | -34.5119 | -18.5952
11.0487 | -34.5123 | -18.6044
11.0548 | -34.5127 | -18.6133
11.0609 | -34.5131 | -18.6226
11.067 | -34.5135 | -18.6318
11.0731 | -34.514 | -18.6408
11.0792 | -34.5144 | -18.6499
11.0853 | -34.5148 | -18.659
11.0914 | -34.5152 | -18.6683
11.0975 | -34.5156 | -18.6774
11.1037 | -34.5161 | -18.6865
11.1098 | -34.5165 | -18.6959
11.1159 | -34.5169 | -18.7047
11.122 | -34.5173 | -18.7138
11.1281 | -34.5178 | -18.723
11.1342 | -34.5182 | -18.7321
11.1403 | -34.5186 | -18.7413
11.1464 | -34.519 | -18.7505
11.1525 | -34.5194 | -18.7595
11.1586 | -34.5198 | -18.7687
11.1647 | -34.5203 | -18.7778
11.1708 | -34.5207 | -18.787
11.1769 | -34.5211 | -18.7962
11.183 | -34.5215 | -18.8052
11.1891 | -34.522 | -18.8142
11.1952 | -34.5224 | -18.8233
11.2013 | -34.5228 | -18.8325
11.2074 | -34.5232 | -18.8416
11.2135 | -34.5237 | -18.8508
11.2196 | -34.5241 | -18.8598
11.2257 | -34.5245 | -18.8689
11.2318 | -34.525 | -18.8779
11.2379 | -34.5254 | -18.8871
11.244 | -34.5259 | -18.8964
11.2502 | -34.5263 | -18.9054
11.2563 | -34.5267 | -18.9147
11.2624 | -34.5272 | -18.9237
11.2685 | -34.5276 | -18.9327
11.2746 | -34.528 | -18.9418
11.2807 | -34.5284 | -18.9512
11.2868 | -34.5289 | -18.9604
11.2929 | -34.5293 | -18.9692
11.299 | -34.5297 | -18.9784
11.3051 | -34.5301 | -18.9877
11.3112 | -34.5305 | -18.9968
11.3173 | -34.531 | -19.0059
11.3234 | -34.5314 | -19.0147
11.3295 | -34.5318 | -19.0239
11.3356 | -34.5323 | -19.0331
11.3417 | -34.5327 | -19.0421
11.3478 | -34.5331 | -19.0512
11.3539 | -34.5336 | -19.0605
11.36 | -34.534 | -19.0694
11.3661 | -34.5344 | -19.0785
11.3722 | -34.5349 | -19.0876
11.3783 | -34.5353 | -19.0967
11.3844 | -34.5358 | -19.1059
11.3906 | -34.5362 | -19.1149
11.3967 | -34.5367 | -19.124
11.4028 | -34.5371 | -19.1332
11.4089 | -34.5375 | -19.1421
11.415 | -34.5379 | -19.1511
11.4211 | -34.5383 | -19.1604
11.4272 | -34.5388 | -19.1695
11.4333 | -34.5392 | -19.1786
11.4394 | -34.5396 | -19.1876
11.4455 | -34.54 | -19.1965
11.4516 | -34.5405 | -19.2056
11.4577 | -34.5409 | -19.2146
11.4638 | -34.5413 | -19.2233
11.4699 | -34.5417 | -19.2324
11.476 | -34.5422 | -19.2417
11.4821 | -34.5426 | -19.2507
11.4882 | -34.543 | -19.2598
11.4943 | -34.5434 | -19.2688
11.5004 | -34.5439 | -19.2776
11.5065 | -34.5443 | -19.2866
11.5126 | -34.5447 | -19.2955
11.5187 | -34.5452 | -19.3046
11.5248 | -34.5456 | -19.3139
11.5309 | -34.546 | -19.3228
11.5371 | -34.5465 | -19.3315
11.5432 | -34.5469 | -19.3405
11.5493 | -34.5473 | -19.3496
11.5554 | -34.5478 | -19.3586
11.5615 | -34.5482 | -19.3678
11.5676 | -34.5486 | -19.3769
11.5737 | -34.5491 | -19.3858
11.5798 | -34.5495 | -19.3948
11.5859 | -34.5499 | -19.4039
11.592 | -34.5504 | -19.4129
11.5981 | -34.5508 | -19.4219
11.6042 | -34.5513 | -19.431
11.6103 | -34.5517 | -19.44
11.6164 | -34.5521 | -19.449
11.6225 | -34.5526 | -19.458
11.6286 | -34.553 | -19.4668
11.6347 | -34.5535 | -19.4761
11.6408 | -34.5539 | -19.4849
11.6469 | -34.5543 | -19.4939
11.653 | -34.5548 | -19.5028
11.6591 | -34.5552 | -19.5117
11.6652 | -34.5556 | -19.5204
11.6713 | -34.5561 | -19.5294
11.6775 | -34.5565 | -19.5382
11.6836 | -34.5569 | -19.5471
11.6897 | -34.5574 | -19.5561
11.6958 | -34.5579 | -19.565
11.7019 | -34.5583 | -19.574
11.708 | -34.5588 | -19.5832
11.7141 | -34.5592 | -19.5922
11.7202 | -34.5597 | -19.601
11.7263 | -34.5601 | -19.6101
11.7324 | -34.5605 | -19.6191
11.7385 | -34.5609 | -19.6281
11.7446 | -34.5613 | -19.6369
11.7507 | -34.5618 | -19.6458
11.7568 | -34.5622 | -19.655
11.7629 | -34.5627 | -19.6642
11.769 | -34.5631 | -19.6731
11.7751 | -34.5636 | -19.6819
11.7812 | -34.564 | -19.6908
11.7873 | -34.5645 | -19.6999
11.7934 | -34.5649 | -19.7091
11.7995 | -34.5654 | -19.7181
11.8056 | -34.5658 | -19.7272
11.8117 | -34.5662 | -19.7361
11.8178 | -34.5667 | -19.745
11.824 | -34.5672 | -19.7538
11.8301 | -34.5676 | -19.7625
11.8362 | -34.568 | -19.7714
11.8423 | -34.5685 | -19.7804
11.8484 | -34.5689 | -19.7892
11.8545 | -34.5694 | -19.7982
11.8606 | -34.5698 | -19.807
11.8667 | -34.5702 | -19.8162
11.8728 | -34.5707 | -19.825
11.8789 | -34.5711 | -19.8339
11.885 | -34.5716 | -19.843
11.8911 | -34.572 | -19.852
11.8972 | -34.5725 | -19.8609
11.9033 | -34.573 | -19.8698
11.9094 | -34.5734 | -19.8788
11.9155 | -34.5738 | -19.8879
11.9216 | -34.5743 | -19.8964
11.9277 | -34.5747 | -19.9053
11.9338 | -34.5751 | -19.9143
11.9399 | -34.5756 | -19.9233
11.946 | -34.576 | -19.9321
11.9521 | -34.5764 | -19.941
11.9582 | -34.5769 | -19.9497
11.9644 | -34.5773 | -19.9585
11.9705 | -34.5778 | -19.9675
11.9766 | -34.5782 | -19.9764
11.9827 | -34.5787 | -19.9852
11.9888 | -34.5791 | -19.9943
11.9949 | -34.5796 | -20.0028
12.001 | -34.58 | -20.0119
12.0071 | -34.5805 | -20.0208
12.0132 | -34.5809 | -20.0296
12.0193 | -34.5813 | -20.0387
12.0254 | -34.5817 | -20.0475
12.0315 | -34.5822 | -20.0565
12.0376 | -34.5826 | -20.0653
12.0437 | -34.5831 | -20.0743
12.0498 | -34.5836 | -20.0834
12.0559 | -34.584 | -20.0921
12.062 | -34.5844 | -20.1005
12.0681 | -34.5849 | -20.1094
12.0742 | -34.5853 | -20.1183
12.0803 | -34.5858 | -20.1272
12.0864 | -34.5862 | -20.1361
12.0925 | -34.5867 | -20.1452
12.0986 | -34.5871 | -20.154
12.1047 | -34.5876 | -20.1628
12.1109 | -34.588 | -20.1716
12.117 | -34.5885 | -20.1804
12.1231 | -34.5889 | -20.1893
12.1292 | -34.5894 | -20.1981
12.1353 | -34.5898 | -20.2068
12.1414 | -34.5903 | -20.2158
12.1475 | -34.5907 | -20.2249
12.1536 | -34.5911 | -20.2338
12.1597 | -34.5916 | -20.2428
12.1658 | -34.592 | -20.2516
12.1719 | -34.5925 | -20.2602
12.178 | -34.5929 | -20.2691
12.1841 | -34.5934 | -20.2781
12.1902 | -34.5938 | -20.2869
12.1963 | -34.5943 | -20.2958
12.2024 | -34.5947 | -20.3044
12.2085 | -34.5952 | -20.3132
12.2146 | -34.5956 | -20.322
12.2207 | -34.5961 | -20.3308
12.2268 | -34.5965 | -20.3397
12.2329 | -34.597 | -20.3486
12.239 | -34.5974 | -20.3571
12.2451 | -34.5979 | -20.366
12.2513 | -34.5983 | -20.3749
12.2574 | -34.5988 | -20.3837
12.2635 | -34.5992 | -20.3926
12.2696 | -34.5997 | -20.4013
12.2757 | -34.6001 | -20.4098
12.2818 | -34.6006 | -20.4186
12.2879 | -34.601 | -20.4275
12.294 | -34.6015 | -20.4361
12.3001 | -34.6019 | -20.4449
12.3062 | -34.6024 | -20.4537
12.3123 | -34.6028 | -20.4621
12.3184 | -34.6032 | -20.471
12.3245 | -34.6037 | -20.4797
12.3306 | -34.6041 | -20.4883
12.3367 | -34.6045 | -20.4973
12.3428 | -34.605 | -20.506
12.3489 | -34.6055 | -20.5145
12.355 | -34.6059 | -20.5234
12.3611 | -34.6064 | -20.5324
12.3672 | -34.6068 | -20.5411
12.3733 | -34.6073 | -20.5499
12.3794 | -34.6077 | -20.5584
12.3855 | -34.6081 | -20.567
12.3916 | -34.6086 | -20.5758
12.3978 | -34.609 | -20.5846
12.4039 | -34.6095 | -20.5933
12.41 | -34.6099 | -20.6021
12.4161 | -34.6104 | -20.6106
12.4222 | -34.6108 | -20.6193
12.4283 | -34.6113 | -20.6281
12.4344 | -34.6118 | -20.637
12.4405 | -34.6122 | -20.6458
12.4466 | -34.6126 | -20.6546
12.4527 | -34.6131 | -20.6634
12.4588 | -34.6135 | -20.6721
12.4649 | -34.614 | -20.6808
12.471 | -34.6145 | -20.6895
12.4771 | -34.6149 | -20.6984
12.4832 | -34.6153 | -20.7071
12.4893 | -34.6158 | -20.7158
12.4954 | -34.6162 | -20.7244
12.5015 | -34.6166 | -20.7331
12.5076 | -34.6171 | -20.742
12.5137 | -34.6176 | -20.7507
12.5198 | -34.618 | -20.7593
12.5259 | -34.6185 | -20.7681
12.532 | -34.6189 | -20.7769
12.5382 | -34.6194 | -20.7857
12.5443 | -34.6198 | -20.7944
12.5504 | -34.6202 | -20.8033
12.5565 | -34.6207 | -20.8119
12.5626 | -34.6212 | -20.8208
12.5687 | -34.6216 | -20.8296
12.5748 | -34.6221 | -20.8383
12.5809 | -34.6225 | -20.8471
12.587 | -34.6229 | -20.856
12.5931 | -34.6234 | -20.8647
12.5992 | -34.6239 | -20.8734
12.6053 | -34.6243 | -20.8821
12.6114 | -34.6248 | -20.8906
12.6175 | -34.6252 | -20.8996
12.6236 | -34.6256 | -20.9082
12.6297 | -34.6261 | -20.9167
12.6358 | -34.6265 | -20.9253
12.6419 | -34.627 | -20.934
12.648 | -34.6274 | -20.9428
12.6541 | -34.6278 | -20.9515
12.6602 | -34.6283 | -20.96
12.6663 | -34.6287 | -20.9687
12.6724 | -34.6292 | -20.9774
12.6785 | -34.6296 | -20.986
12.6847 | -34.6301 | -20.9947
12.6908 | -34.6305 | -21.0034
12.6969 | -34.631 | -21.012
12.703 | -34.6314 | -21.0206
12.7091 | -34.6318 | -21.0294
12.7152 | -34.6323 | -21.0379
12.7213 | -34.6327 | -21.0466
12.7274 | -34.6332 | -21.0554
12.7335 | -34.6336 | -21.0639
12.7396 | -34.634 | -21.0725
12.7457 | -34.6345 | -21.0812
12.7518 | -34.6349 | -21.0899
12.7579 | -34.6354 | -21.0987
12.764 | -34.6358 | -21.1075
12.7701 | -34.6363 | -21.1162
12.7762 | -34.6367 | -21.1249
12.7823 | -34.6372 | -21.1336
12.7884 | -34.6376 | -21.1424
12.7945 | -34.6381 | -21.1512
12.8006 | -34.6386 | -21.1597
12.8067 | -34.639 | -21.1684
12.8128 | -34.6394 | -21.1769
12.8189 | -34.6399 | -21.1856
12.8251 | -34.6403 | -21.1943
12.8312 | -34.6408 | -21.2032
12.8373 | -34.6412 | -21.2118
12.8434 | -34.6417 | -21.2206
12.8495 | -34.6421 | -21.2292
12.8556 | -34.6425 | -21.2377
12.8617 | -34.643 | -21.2464
12.8678 | -34.6435 | -21.2552
12.8739 | -34.6439 | -21.2637
12.88 | -34.6443 | -21.2723
12.8861 | -34.6448 | -21.2809
12.8922 | -34.6452 | -21.2896
12.8983 | -34.6457 | -21.2983
12.9044 | -34.6461 | -21.3068
12.9105 | -34.6466 | -21.3153
12.9166 | -34.647 | -21.3239
12.9227 | -34.6475 | -21.3326
12.9288 | -34.6479 | -21.3414
12.9349 | -34.6484 | -21.35
12.941 | -34.6488 | -21.3585
12.9471 | -34.6493 | -21.3673
12.9532 | -34.6497 | -21.3759
12.9593 | -34.6502 | -21.3848
12.9654 | -34.6506 | -21.3935
12.9716 | -34.6511 | -21.4023
12.9777 | -34.6515 | -21.4109
12.9838 | -34.652 | -21.4195
12.9899 | -34.6524 | -21.4281
12.996 | -34.6529 | -21.4368
13.0021 | -34.6533 | -21.4455
13.0082 | -34.6538 | -21.454
13.0143 | -34.6543 | -21.4625
13.0204 | -34.6547 | -21.4711
13.0265 | -34.6551 | -21.4797
13.0326 | -34.6556 | -21.4883
13.0387 | -34.656 | -21.497
13.0448 | -34.6565 | -21.5056
13.0509 | -34.6569 | -21.5142
13.057 | -34.6574 | -21.5227
13.0631 | -34.6578 | -21.5313
13.0692 | -34.6583 | -21.5398
13.0753 | -34.6587 | -21.5485
13.0814 | -34.6592 | -21.5573
13.0875 | -34.6596 | -21.5657
13.0936 | -34.6601 | -21.5744
13.0997 | -34.6605 | -21.583
13.1058 | -34.661 | -21.5915
13.112 | -34.6614 | -21.6002
13.1181 | -34.6619 | -21.6087
13.1242 | -34.6623 | -21.6174
13.1303 | -34.6628 | -21.626
13.1364 | -34.6632 | -21.6345
13.1425 | -34.6636 | -21.6431
13.1486 | -34.6641 | -21.6517
13.1547 | -34.6645 | -21.6602
13.1608 | -34.665 | -21.6688
13.1669 | -34.6654 | -21.6773
13.173 | -34.6659 | -21.686
13.1791 | -34.6663 | -21.6945
13.1852 | -34.6667 | -21.7031
13.1913 | -34.6672 | -21.7115
13.1974 | -34.6676 | -21.7202
13.2035 | -34.668 | -21.7288
13.2096 | -34.6685 | -21.7373
13.2157 | -34.6689 | -21.7458
13.2218 | -34.6694 | -21.7544
13.2279 | -34.6698 | -21.763
13.234 | -34.6703 | -21.7716
13.2401 | -34.6707 | -21.7802
13.2462 | -34.6712 | -21.7888
13.2524 | -34.6716 | -21.7975
13.2585 | -34.6721 | -21.8059
13.2646 | -34.6725 | -21.8145
13.2707 | -34.673 | -21.8232
13.2768 | -34.6734 | -21.8319
13.2829 | -34.6739 | -21.8405
13.289 | -34.6743 | -21.849
13.2951 | -34.6748 | -21.8575
13.3012 | -34.6752 | -21.8661
13.3073 | -34.6756 | -21.8748
13.3134 | -34.6761 | -21.8834
13.3195 | -34.6765 | -21.892
13.3256 | -34.6769 | -21.9004
13.3317 | -34.6774 | -21.9089
13.3378 | -34.6778 | -21.9175
13.3439 | -34.6783 | -21.9262
13.35 | -34.6787 | -21.9348
13.3561 | -34.6791 | -21.9433
13.3622 | -34.6795 | -21.9517
13.3683 | -34.68 | -21.9602
13.3744 | -34.6804 | -21.9687
13.3805 | -34.6809 | -21.9774
13.3866 | -34.6813 | -21.9859
13.3927 | -34.6817 | -21.9944
13.3989 | -34.6822 | -22.0029
13.405 | -34.6826 | -22.0115
13.4111 | -34.683 | -22.0199
13.4172 | -34.6835 | -22.0285
13.4233 | -34.6839 | -22.0372
13.4294 | -34.6843 | -22.0458
13.4355 | -34.6848 | -22.0541
13.4416 | -34.6852 | -22.0627
13.4477 | -34.6857 | -22.0713
13.4538 | -34.6861 | -22.0799
13.4599 | -34.6866 | -22.0884
13.466 | -34.687 | -22.097
13.4721 | -34.6875 | -22.1056
13.4782 | -34.6879 | -22.1141
13.4843 | -34.6884 | -22.1227
13.4904 | -34.6888 | -22.1313
13.4965 | -34.6892 | -22.1399
13.5026 | -34.6897 | -22.1484
13.5087 | -34.6901 | -22.1568
13.5148 | -34.6906 | -22.1654
13.5209 | -34.691 | -22.1741
13.527 | -34.6914 | -22.1826
13.5331 | -34.6919 | -22.1911
13.5393 | -34.6923 | -22.1995
13.5454 | -34.6928 | -22.208
13.5515 | -34.6932 | -22.2166
13.5576 | -34.6937 | -22.2251
13.5637 | -34.6941 | -22.2338
13.5698 | -34.6946 | -22.2424
13.5759 | -34.695 | -22.2506
13.582 | -34.6955 | -22.259
13.5881 | -34.6959 | -22.2675
13.5942 | -34.6964 | -22.2761
13.6003 | -34.6968 | -22.2846
13.6064 | -34.6973 | -22.2933
13.6125 | -34.6977 | -22.3018
13.6186 | -34.6982 | -22.3102
13.6247 | -34.6986 | -22.3188
13.6308 | -34.6991 | -22.3273
13.6369 | -34.6995 | -22.3359
13.643 | -34.6999 | -22.3446
13.6491 | -34.7004 | -22.3529
13.6552 | -34.7008 | -22.3613
13.6613 | -34.7013 | -22.3698
13.6674 | -34.7017 | -22.3785
13.6735 | -34.7021 | -22.3871
13.6796 | -34.7026 | -22.3956
13.6858 | -34.703 | -22.4041
13.6919 | -34.7035 | -22.4126
13.698 | -34.7039 | -22.421
13.7041 | -34.7044 | -22.4296
13.7102 | -34.7048 | -22.438
13.7163 | -34.7053 | -22.4465
13.7224 | -34.7057 | -22.455
13.7285 | -34.7062 | -22.4635
13.7346 | -34.7066 | -22.472
13.7407 | -34.707 | -22.4804
13.7468 | -34.7075 | -22.4889
13.7529 | -34.7079 | -22.4973
13.759 | -34.7084 | -22.5059
13.7651 | -34.7088 | -22.5144
13.7712 | -34.7093 | -22.5229
13.7773 | -34.7097 | -22.5313
13.7834 | -34.7102 | -22.5398
13.7895 | -34.7107 | -22.5484
13.7956 | -34.7111 | -22.557
13.8017 | -34.7115 | -22.5654
13.8078 | -34.712 | -22.5738
13.8139 | -34.7124 | -22.5822
13.82 | -34.7128 | -22.5906
13.8262 | -34.7133 | -22.5991
13.8323 | -34.7137 | -22.6076
13.8384 | -34.7142 | -22.6162
13.8445 | -34.7146 | -22.6247
13.8506 | -34.7151 | -22.6333
13.8567 | -34.7155 | -22.6418
13.8628 | -34.716 | -22.6503
13.8689 | -34.7164 | -22.6588
13.875 | -34.7169 | -22.6673
13.8811 | -34.7173 | -22.6756
13.8872 | -34.7177 | -22.684
13.8933 | -34.7182 | -22.6923
13.8994 | -34.7186 | -22.7009
13.9055 | -34.7191 | -22.7094
13.9116 | -34.7195 | -22.7179
13.9177 | -34.7199 | -22.7264
13.9238 | -34.7204 | -22.7348
13.9299 | -34.7208 | -22.7432
13.936 | -34.7213 | -22.7517
13.9421 | -34.7217 | -22.7602
13.9482 | -34.7222 | -22.7686
13.9543 | -34.7226 | -22.7771
13.9604 | -34.723 | -22.7855
13.9665 | -34.7235 | -22.7938
13.9727 | -34.724 | -22.8024
13.9788 | -34.7244 | -22.8108
13.9849 | -34.7249 | -22.8192
13.991 | -34.7253 | -22.8277
13.9971 | -34.7258 | -22.836
14.0032 | -34.7262 | -22.8444
14.0093 | -34.7267 | -22.8529
14.0154 | -34.7271 | -22.8613
14.0215 | -34.7275 | -22.8696
14.0276 | -34.728 | -22.8779
14.0337 | -34.7285 | -22.8863
14.0398 | -34.7289 | -22.8947
14.0459 | -34.7294 | -22.903
14.052 | -34.7298 | -22.9115
14.0581 | -34.7302 | -22.92
14.0642 | -34.7306 | -22.9284
14.0703 | -34.7311 | -22.9368
14.0764 | -34.7316 | -22.945
14.0825 | -34.732 | -22.9535
14.0886 | -34.7325 | -22.9619
14.0947 | -34.7329 | -22.9703
14.1008 | -34.7334 | -22.9787
14.1069 | -34.7338 | -22.9869
14.1131 | -34.7343 | -22.9953
14.1192 | -34.7347 | -23.0036
14.1253 | -34.7352 | -23.0119
14.1314 | -34.7356 | -23.0204
14.1375 | -34.7361 | -23.0287
14.1436 | -34.7366 | -23.0371
14.1497 | -34.737 | -23.0455
14.1558 | -34.7375 | -23.0539
14.1619 | -34.7379 | -23.0623
14.168 | -34.7384 | -23.0708
14.1741 | -34.7388 | -23.079
14.1802 | -34.7393 | -23.0873
14.1863 | -34.7397 | -23.0958
14.1924 | -34.7402 | -23.1042
14.1985 | -34.7407 | -23.1126
14.2046 | -34.7411 | -23.121
14.2107 | -34.7416 | -23.1291
14.2168 | -34.742 | -23.1375
14.2229 | -34.7425 | -23.1459
14.229 | -34.7429 | -23.1543
14.2351 | -34.7434 | -23.1626
14.2412 | -34.7438 | -23.171
14.2473 | -34.7443 | -23.1793
14.2534 | -34.7448 | -23.1876
14.2596 | -34.7452 | -23.196
14.2657 | -34.7456 | -23.2043
14.2718 | -34.7461 | -23.2127
14.2779 | -34.7465 | -23.2211
14.284 | -34.747 | -23.2294
14.2901 | -34.7475 | -23.2378
14.2962 | -34.7479 | -23.2462
14.3023 | -34.7484 | -23.2546
14.3084 | -34.7488 | -23.2629
14.3145 | -34.7493 | -23.2712
14.3206 | -34.7498 | -23.2797
14.3267 | -34.7502 | -23.2881
14.3328 | -34.7507 | -23.2964
14.3389 | -34.7511 | -23.3047
14.345 | -34.7516 | -23.313
14.3511 | -34.752 | -23.3213
14.3572 | -34.7525 | -23.3297
14.3633 | -34.753 | -23.3381
14.3694 | -34.7534 | -23.3464
14.3755 | -34.7539 | -23.3547
14.3816 | -34.7544 | -23.3632
14.3877 | -34.7549 | -23.3714
14.3938 | -34.7553 | -23.3799
14.4 | -34.7558 | -23.3882
14.4061 | -34.7562 | -23.3965
14.4122 | -34.7567 | -23.4048
14.4183 | -34.7571 | -23.413
14.4244 | -34.7576 | -23.4214
14.4305 | -34.7581 | -23.4297
14.4366 | -34.7585 | -23.4381
14.4427 | -34.759 | -23.4465
14.4488 | -34.7595 | -23.4549
14.4549 | -34.7599 | -23.463
14.461 | -34.7604 | -23.4713
14.4671 | -34.7609 | -23.4797
14.4732 | -34.7613 | -23.488
14.4793 | -34.7618 | -23.4964
14.4854 | -34.7622 | -23.5046
14.4915 | -34.7627 | -23.5128
14.4976 | -34.7632 | -23.521
14.5037 | -34.7636 | -23.5293
14.5098 | -34.7641 | -23.5376
14.5159 | -34.7645 | -23.5459
14.522 | -34.765 | -23.5542
14.5281 | -34.7655 | -23.5624
14.5342 | -34.7659 | -23.5708
14.5403 | -34.7664 | -23.579
14.5465 | -34.7669 | -23.5873
14.5526 | -34.7673 | -23.5957
14.5587 | -34.7678 | -23.604
14.5648 | -34.7683 | -23.6122
14.5709 | -34.7687 | -23.6205
14.577 | -34.7692 | -23.6287
14.5831 | -34.7696 | -23.6371
14.5892 | -34.7701 | -23.6453
14.5953 | -34.7706 | -23.6534
14.6014 | -34.771 | -23.6616
14.6075 | -34.7715 | -23.6698
14.6136 | -34.772 | -23.6781
14.6197 | -34.7724 | -23.6865
14.6258 | -34.7729 | -23.6948
14.6319 | -34.7734 | -23.703
14.638 | -34.7738 | -23.7113
14.6441 | -34.7743 | -23.7195
14.6502 | -34.7747 | -23.7279
14.6563 | -34.7752 | -23.7361
14.6624 | -34.7757 | -23.7444
14.6685 | -34.7762 | -23.7523
14.6746 | -34.7766 | -23.7607
14.6807 | -34.7771 | -23.7689
14.6869 | -34.7775 | -23.7771
14.693 | -34.778 | -23.7853
14.6991 | -34.7785 | -23.7934
14.7052 | -34.779 | -23.8015
14.7113 | -34.7795 | -23.8099
14.7174 | -34.7799 | -23.8182
14.7235 | -34.7804 | -23.8264
14.7296 | -34.7809 | -23.8346
14.7357 | -34.7814 | -23.8427
14.7418 | -34.7818 | -23.851
14.7479 | -34.7823 | -23.8592
14.754 | -34.7828 | -23.8674
14.7601 | -34.7832 | -23.8755
14.7662 | -34.7837 | -23.8836
14.7723 | -34.7842 | -23.8915
14.7784 | -34.7846 | -23.8998
14.7845 | -34.7851 | -23.908
14.7906 | -34.7856 | -23.9164
14.7967 | -34.786 | -23.9246
14.8028 | -34.7865 | -23.9328
14.8089 | -34.787 | -23.9409
14.815 | -34.7875 | -23.9492
14.8211 | -34.7879 | -23.9575
14.8272 | -34.7884 | -23.9658
14.8334 | -34.7889 | -23.974
14.8395 | -34.7893 | -23.9821
14.8456 | -34.7898 | -23.9901
14.8517 | -34.7903 | -23.9983
14.8578 | -34.7908 | -24.0066
14.8639 | -34.7912 | -24.0148
14.87 | -34.7917 | -24.023
14.8761 | -34.7922 | -24.0311
14.8822 | -34.7927 | -24.0393
14.8883 | -34.7931 | -24.0475
14.8944 | -34.7936 | -24.0557
14.9005 | -34.7941 | -24.0639
14.9066 | -34.7945 | -24.0721
14.9127 | -34.795 | -24.0801
14.9188 | -34.7955 | -24.0883
14.9249 | -34.796 | -24.0965
14.931 | -34.7964 | -24.1047
14.9371 | -34.7969 | -24.1128
14.9432 | -34.7974 | -24.1211
14.9493 | -34.7979 | -24.1291
14.9554 | -34.7983 | -24.1372
14.9615 | -34.7988 | -24.1454
14.9676 | -34.7993 | -24.1537
14.9738 | -34.7997 | -24.1617
14.9799 | -34.8002 | -24.1701
14.986 | -34.8007 | -24.1782
14.9921 | -34.8012 | -24.1864
14.9982 | -34.8016 | -24.1945
15.0043 | -34.8021 | -24.2026
15.0104 | -34.8026 | -24.2109
15.0165 | -34.8031 | -24.2192
15.0226 | -34.8036 | -24.2271
15.0287 | -34.804 | -24.2354
15.0348 | -34.8045 | -24.2435
15.0409 | -34.805 | -24.2518
15.047 | -34.8055 | -24.26
15.0531 | -34.806 | -24.268
15.0592 | -34.8065 | -24.2762
15.0653 | -34.807 | -24.2845
15.0714 | -34.8074 | -24.2925
15.0775 | -34.8079 | -24.3006
15.0836 | -34.8084 | -24.3087
15.0897 | -34.8089 | -24.3167
15.0958 | -34.8094 | -24.3248
15.1019 | -34.8099 | -24.3329
15.108 | -34.8104 | -24.3411
15.1141 | -34.8109 | -24.3492
15.1203 | -34.8114 | -24.3574
15.1264 | -34.8118 | -24.3655
15.1325 | -34.8123 | -24.3737
15.1386 | -34.8128 | -24.3817
15.1447 | -34.8133 | -24.3898
15.1508 | -34.8138 | -24.398
15.1569 | -34.8143 | -24.4061
15.163 | -34.8148 | -24.414
15.1691 | -34.8153 | -24.4221
15.1752 | -34.8157 | -24.4302
15.1813 | -34.8162 | -24.4383
15.1874 | -34.8167 | -24.4464
15.1935 | -34.8172 | -24.4543
15.1996 | -34.8177 | -24.4623
15.2057 | -34.8181 | -24.4704
15.2118 | -34.8186 | -24.4785
15.2179 | -34.8191 | -24.4866
15.224 | -34.8196 | -24.4947
15.2301 | -34.8201 | -24.5025
15.2362 | -34.8206 | -24.5106
15.2423 | -34.821 | -24.5187
15.2484 | -34.8215 | -24.5269
15.2545 | -34.822 | -24.5349
15.2607 | -34.8225 | -24.5429
15.2668 | -34.823 | -24.5506
15.2729 | -34.8235 | -24.5586
15.279 | -34.824 | -24.5668
15.2851 | -34.8245 | -24.5748
15.2912 | -34.825 | -24.583
15.2973 | -34.8255 | -24.5912
15.3034 | -34.826 | -24.5992
15.3095 | -34.8265 | -24.6073
15.3156 | -34.827 | -24.6152
15.3217 | -34.8275 | -24.6233
15.3278 | -34.8279 | -24.6315
15.3339 | -34.8285 | -24.6394
15.34 | -34.8289 | -24.6473
15.3461 | -34.8294 | -24.6552
15.3522 | -34.8299 | -24.6633
15.3583 | -34.8304 | -24.6714
15.3644 | -34.8309 | -24.6794
15.3705 | -34.8314 | -24.6872
15.3766 | -34.8319 | -24.6953
15.3827 | -34.8324 | -24.7032
15.3888 | -34.8329 | -24.7113
15.3949 | -34.8333 | -24.7194
15.401 | -34.8338 | -24.7275
15.4072 | -34.8343 | -24.7352
15.4133 | -34.8348 | -24.7433
15.4194 | -34.8353 | -24.7513
15.4255 | -34.8358 | -24.7594
15.4316 | -34.8362 | -24.7674
15.4377 | -34.8367 | -24.7756
15.4438 | -34.8373 | -24.7834
15.4499 | -34.8378 | -24.7914
15.456 | -34.8383 | -24.7993
15.4621 | -34.8387 | -24.8072
15.4682 | -34.8392 | -24.8154
15.4743 | -34.8397 | -24.8232
15.4804 | -34.8402 | -24.8313
15.4865 | -34.8407 | -24.8393
15.4926 | -34.8412 | -24.8474
15.4987 | -34.8417 | -24.8555
15.5048 | -34.8422 | -24.8635
15.5109 | -34.8427 | -24.8714
15.517 | -34.8432 | -24.8794
15.5231 | -34.8437 | -24.8873
15.5292 | -34.8442 | -24.8953
15.5353 | -34.8447 | -24.9034
15.5414 | -34.8452 | -24.9113
15.5476 | -34.8457 | -24.9189
15.5537 | -34.8462 | -24.9269
15.5598 | -34.8467 | -24.935
15.5659 | -34.8472 | -24.943
15.572 | -34.8477 | -24.9509
15.5781 | -34.8482 | -24.9588
15.5842 | -34.8487 | -24.9666
15.5903 | -34.8492 | -24.9746
15.5964 | -34.8497 | -24.9825
15.6025 | -34.8502 | -24.9904
15.6086 | -34.8507 | -24.9983
15.6147 | -34.8512 | -25.0064
15.6208 | -34.8517 | -25.0142
15.6269 | -34.8522 | -25.0221
15.633 | -34.8527 | -25.0302
15.6391 | -34.8532 | -25.0382
15.6452 | -34.8537 | -25.0463
15.6513 | -34.8542 | -25.0541
15.6574 | -34.8547 | -25.0621
15.6635 | -34.8552 | -25.07
15.6696 | -34.8557 | -25.078
15.6757 | -34.8562 | -25.086
15.6818 | -34.8567 | -25.094
15.688 | -34.8572 | -25.1019
15.6941 | -34.8577 | -25.1098
15.7002 | -34.8582 | -25.1177
15.7063 | -34.8587 | -25.1255
15.7124 | -34.8592 | -25.1336
15.7185 | -34.8597 | -25.1415
15.7246 | -34.8602 | -25.1494
15.7307 | -34.8607 | -25.1572
15.7368 | -34.8612 | -25.1652
15.7429 | -34.8617 | -25.173
15.749 | -34.8622 | -25.181
15.7551 | -34.8627 | -25.189
15.7612 | -34.8633 | -25.1966
15.7673 | -34.8638 | -25.2044
15.7734 | -34.8643 | -25.2123
15.7795 | -34.8648 | -25.2202
15.7856 | -34.8653 | -25.2279
15.7917 | -34.8658 | -25.2355
15.7978 | -34.8663 | -25.2436
15.8039 | -34.8667 | -25.2515
15.81 | -34.8672 | -25.2594
15.8161 | -34.8677 | -25.2672
15.8222 | -34.8682 | -25.275
15.8283 | -34.8687 | -25.2826
15.8345 | -34.8693 | -25.2906
15.8406 | -34.8698 | -25.2984
15.8467 | -34.8703 | -25.3063
15.8528 | -34.8708 | -25.3143
15.8589 | -34.8713 | -25.3222
15.865 | -34.8718 | -25.3298
15.8711 | -34.8723 | -25.3377
15.8772 | -34.8728 | -25.3455
15.8833 | -34.8734 | -25.3534
15.8894 | -34.8739 | -25.3613
15.8955 | -34.8744 | -25.3691
15.9016 | -34.8749 | -25.3767
15.9077 | -34.8754 | -25.3846
15.9138 | -34.8759 | -25.3925
15.9199 | -34.8764 | -25.4003
15.926 | -34.8769 | -25.4081
15.9321 | -34.8774 | -25.4157
15.9382 | -34.8779 | -25.4236
15.9443 | -34.8784 | -25.4314
15.9504 | -34.8789 | -25.4393
15.9565 | -34.8794 | -25.4472
15.9626 | -34.8799 | -25.4552
15.9687 | -34.8804 | -25.4628
15.9749 | -34.8809 | -25.4706
15.981 | -34.8814 | -25.4783
15.9871 | -34.8819 | -25.4862
15.9932 | -34.8824 | -25.4938
15.9993 | -34.8829 | -25.5016
16.0054 | -34.8834 | -25.5091
16.0115 | -34.8839 | -25.517
16.0176 | -34.8844 | -25.525
16.0237 | -34.8849 | -25.5328
16.0298 | -34.8854 | -25.5405
16.0359 | -34.8859 | -25.5483
16.042 | -34.8864 | -25.556
16.0481 | -34.8869 | -25.5638
16.0542 | -34.8875 | -25.5717
16.0603 | -34.888 | -25.5794
16.0664 | -34.8885 | -25.5873
16.0725 | -34.889 | -25.5949
16.0786 | -34.8895 | -25.6025
16.0847 | -34.89 | -25.6105
16.0908 | -34.8905 | -25.6184
16.0969 | -34.891 | -25.6263
16.103 | -34.8915 | -25.6341
16.1091 | -34.892 | -25.6417
16.1152 | -34.8925 | -25.6495
16.1214 | -34.893 | -25.6573
16.1275 | -34.8935 | -25.6651
16.1336 | -34.894 | -25.6728
16.1397 | -34.8945 | -25.6805
16.1458 | -34.895 | -25.6882
16.1519 | -34.8955 | -25.6961
16.158 | -34.896 | -25.7038
16.1641 | -34.8965 | -25.7115
16.1702 | -34.897 | -25.7194
16.1763 | -34.8975 | -25.7272
16.1824 | -34.8981 | -25.7347
16.1885 | -34.8986 | -25.7425
16.1946 | -34.8991 | -25.7503
16.2007 | -34.8996 | -25.758
16.2068 | -34.9001 | -25.7658
16.2129 | -34.9006 | -25.7737
16.219 | -34.9011 | -25.7812
16.2251 | -34.9016 | -25.7888
16.2312 | -34.9022 | -25.7964
16.2373 | -34.9027 | -25.804
16.2434 | -34.9032 | -25.8118
16.2495 | -34.9037 | -25.8194
16.2556 | -34.9042 | -25.8271
16.2618 | -34.9048 | -25.8349
16.2679 | -34.9053 | -25.8428
16.274 | -34.9057 | -25.8505
16.2801 | -34.9063 | -25.8583
16.2862 | -34.9068 | -25.8659
16.2923 | -34.9073 | -25.8737
16.2984 | -34.9078 | -25.8815
16.3045 | -34.9083 | -25.8893
16.3106 | -34.9088 | -25.8971
16.3167 | -34.9093 | -25.9048
16.3228 | -34.9098 | -25.9124
16.3289 | -34.9103 | -25.9202
16.335 | -34.9108 | -25.9279
16.3411 | -34.9113 | -25.9355
16.3472 | -34.9119 | -25.9434
16.3533 | -34.9124 | -25.9511
16.3594 | -34.9129 | -25.9585
16.3655 | -34.9134 | -25.9662
16.3716 | -34.9138 | -25.9739
16.3777 | -34.9144 | -25.9814
16.3838 | -34.9148 | -25.9892
16.3899 | -34.9154 | -25.9966
16.396 | -34.9159 | -26.0044
16.4021 | -34.9164 | -26.0121
16.4083 | -34.9169 | -26.0198
16.4144 | -34.9174 | -26.0276
16.4205 | -34.9179 | -26.0354
16.4266 | -34.9184 | -26.043
16.4327 | -34.9189 | -26.0506
16.4388 | -34.9194 | -26.0583
16.4449 | -34.92 | -26.0659
16.451 | -34.9205 | -26.0739
16.4571 | -34.921 | -26.0816
16.4632 | -34.9215 | -26.0891
16.4693 | -34.922 | -26.0968
16.4754 | -34.9225 | -26.1046
16.4815 | -34.923 | -26.1123
16.4876 | -34.9234 | -26.1199
16.4937 | -34.9239 | -26.1275
16.4998 | -34.9245 | -26.135
16.5059 | -34.925 | -26.1425
16.512 | -34.9255 | -26.1502
16.5181 | -34.926 | -26.158
16.5242 | -34.9265 | -26.166
16.5303 | -34.927 | -26.1736
16.5364 | -34.9275 | -26.181
16.5425 | -34.928 | -26.1887
16.5487 | -34.9285 | -26.1964
16.5548 | -34.929 | -26.204
16.5609 | -34.9296 | -26.2118
16.567 | -34.9301 | -26.2193
16.5731 | -34.9306 | -26.2271
16.5792 | -34.9311 | -26.2349
16.5853 | -34.9316 | -26.2425
16.5914 | -34.9321 | -26.2502
16.5975 | -34.9326 | -26.258
16.6036 | -34.9331 | -26.2656
16.6097 | -34.9336 | -26.2733
16.6158 | -34.9341 | -26.2809
16.6219 | -34.9346 | -26.2886
16.628 | -34.9351 | -26.2963
16.6341 | -34.9356 | -26.3037
16.6402 | -34.9361 | -26.3111
16.6463 | -34.9366 | -26.3187
16.6524 | -34.9371 | -26.3263
16.6585 | -34.9376 | -26.334
16.6646 | -34.9381 | -26.3416
16.6707 | -34.9387 | -26.3491
16.6768 | -34.9392 | -26.3568
16.6829 | -34.9396 | -26.3643
16.689 | -34.9402 | -26.3721
16.6952 | -34.9407 | -26.3797
16.7013 | -34.9412 | -26.3873
16.7074 | -34.9417 | -26.3947
16.7135 | -34.9422 | -26.4022
16.7196 | -34.9427 | -26.4097
16.7257 | -34.9432 | -26.4172
16.7318 | -34.9437 | -26.4249
16.7379 | -34.9442 | -26.4325
16.744 | -34.9447 | -26.4398
16.7501 | -34.9452 | -26.4473
16.7562 | -34.9457 | -26.455
16.7623 | -34.9462 | -26.4623
16.7684 | -34.9467 | -26.47
16.7745 | -34.9472 | -26.4775
16.7806 | -34.9477 | -26.4848
16.7867 | -34.9482 | -26.4925
16.7928 | -34.9487 | -26.5001
16.7989 | -34.9492 | -26.5077
16.805 | -34.9497 | -26.5153
16.8111 | -34.9502 | -26.5227
16.8172 | -34.9507 | -26.5305
16.8233 | -34.9512 | -26.5381
16.8294 | -34.9517 | -26.5457
16.8356 | -34.9522 | -26.5534
16.8417 | -34.9528 | -26.5608
16.8478 | -34.9533 | -26.5682
16.8539 | -34.9538 | -26.5757
16.86 | -34.9543 | -26.5835
16.8661 | -34.9548 | -26.5909
16.8722 | -34.9553 | -26.5986
16.8783 | -34.9558 | -26.6063
16.8844 | -34.9563 | -26.6136
16.8905 | -34.9568 | -26.6212
16.8966 | -34.9573 | -26.6287
16.9027 | -34.9578 | -26.6362
16.9088 | -34.9583 | -26.6438
16.9149 | -34.9588 | -26.6513
16.921 | -34.9593 | -26.6585
16.9271 | -34.9598 | -26.6661
16.9332 | -34.9603 | -26.6736
16.9393 | -34.9608 | -26.6812
16.9454 | -34.9613 | -26.6889
16.9515 | -34.9618 | -26.6966
16.9576 | -34.9623 | -26.7039
16.9637 | -34.9628 | -26.7114
16.9698 | -34.9633 | -26.7189
16.9759 | -34.9638 | -26.7266
16.9821 | -34.9644 | -26.7341
16.9882 | -34.9649 | -26.7415
16.9943 | -34.9654 | -26.749
17.0004 | -34.9659 | -26.7565
17.0065 | -34.9664 | -26.7641
17.0126 | -34.9669 | -26.7714
17.0187 | -34.9674 | -26.779
17.0248 | -34.9679 | -26.7863
17.0309 | -34.9684 | -26.794
17.037 | -34.9689 | -26.8013
17.0431 | -34.9694 | -26.8087
17.0492 | -34.9699 | -26.8164
17.0553 | -34.9704 | -26.8238
17.0614 | -34.9709 | -26.8313
17.0675 | -34.9715 | -26.8388
17.0736 | -34.972 | -26.8465
17.0797 | -34.9725 | -26.8539
17.0858 | -34.973 | -26.8615
17.0919 | -34.9735 | -26.8689
17.098 | -34.974 | -26.8762
17.1041 | -34.9745 | -26.8834
17.1102 | -34.975 | -26.891
17.1163 | -34.9755 | -26.8984
17.1225 | -34.976 | -26.9057
17.1286 | -34.9766 | -26.913
17.1347 | -34.9771 | -26.9206
17.1408 | -34.9776 | -26.9284
17.1469 | -34.9781 | -26.9359
17.153 | -34.9786 | -26.9434
17.1591 | -34.9791 | -26.9509
17.1652 | -34.9796 | -26.9582
17.1713 | -34.9801 | -26.9656
17.1774 | -34.9806 | -26.9731
17.1835 | -34.9812 | -26.9804
17.1896 | -34.9817 | -26.9878
17.1957 | -34.9822 | -26.9952
17.2018 | -34.9827 | -27.0024
17.2079 | -34.9832 | -27.01
17.214 | -34.9837 | -27.0176
17.2201 | -34.9842 | -27.0251
17.2262 | -34.9847 | -27.0324
17.2323 | -34.9852 | -27.04
17.2384 | -34.9857 | -27.0475
17.2445 | -34.9862 | -27.0549
17.2506 | -34.9867 | -27.0624
17.2567 | -34.9872 | -27.0698
17.2628 | -34.9877 | -27.0772
17.269 | -34.9883 | -27.0847
17.2751 | -34.9888 | -27.0921
17.2812 | -34.9893 | -27.0996
17.2873 | -34.9898 | -27.1072
17.2934 | -34.9903 | -27.1147
17.2995 | -34.9908 | -27.1222
17.3056 | -34.9913 | -27.1295
17.3117 | -34.9918 | -27.137
17.3178 | -34.9923 | -27.1444
17.3239 | -34.9928 | -27.1518
17.33 | -34.9933 | -27.1593
17.3361 | -34.9938 | -27.1669
17.3422 | -34.9943 | -27.1744
17.3483 | -34.9948 | -27.1819
17.3544 | -34.9953 | -27.1896
17.3605 | -34.9958 | -27.1969
17.3666 | -34.9963 | -27.2044
17.3727 | -34.9968 | -27.2118
17.3788 | -34.9973 | -27.2191
17.3849 | -34.9978 | -27.2265
17.391 | -34.9983 | -27.2339
17.3971 | -34.9988 | -27.2414
17.4032 | -34.9993 | -27.2489
17.4094 | -34.9998 | -27.2561
17.4155 | -35.0003 | -27.2636
17.4216 | -35.0008 | -27.2709
17.4277 | -35.0013 | -27.2784
17.4338 | -35.0018 | -27.2859
17.4399 | -35.0023 | -27.2933
17.446 | -35.0029 | -27.3005
17.4521 | -35.0034 | -27.3081
17.4582 | -35.0039 | -27.3155
17.4643 | -35.0043 | -27.323
17.4704 | -35.0048 | -27.3303
17.4765 | -35.0054 | -27.3378
17.4826 | -35.0059 | -27.3451
17.4887 | -35.0064 | -27.3526
17.4948 | -35.0069 | -27.3601
17.5009 | -35.0074 | -27.3676
17.507 | -35.0079 | -27.375
17.5131 | -35.0084 | -27.3825
17.5192 | -35.0088 | -27.3898
17.5253 | -35.0093 | -27.3972
17.5314 | -35.0098 | -27.4045
17.5375 | -35.0103 | -27.4119
17.5436 | -35.0108 | -27.4196
17.5497 | -35.0113 | -27.4269
17.5559 | -35.0118 | -27.4341
17.562 | -35.0123 | -27.4414
17.5681 | -35.0128 | -27.4486
17.5742 | -35.0133 | -27.4562
17.5803 | -35.0138 | -27.4636
17.5864 | -35.0143 | -27.4709
17.5925 | -35.0148 | -27.4783
17.5986 | -35.0153 | -27.4856
17.6047 | -35.0158 | -27.4932
17.6108 | -35.0163 | -27.5004
17.6169 | -35.0168 | -27.5078
17.623 | -35.0173 | -27.5151
17.6291 | -35.0178 | -27.5224
17.6352 | -35.0183 | -27.5299
17.6413 | -35.0188 | -27.5373
17.6474 | -35.0193 | -27.5448
17.6535 | -35.0198 | -27.5522
17.6596 | -35.0203 | -27.5594
17.6657 | -35.0208 | -27.5668
17.6718 | -35.0213 | -27.5741
17.6779 | -35.0218 | -27.5815
17.684 | -35.0223 | -27.5888
17.6901 | -35.0228 | -27.5961
17.6963 | -35.0233 | -27.6034
17.7024 | -35.0238 | -27.6106
17.7085 | -35.0243 | -27.618
17.7146 | -35.0248 | -27.6253
17.7207 | -35.0252 | -27.6327
17.7268 | -35.0258 | -27.6399
17.7329 | -35.0263 | -27.6472
17.739 | -35.0268 | -27.6546
17.7451 | -35.0273 | -27.6618
17.7512 | -35.0278 | -27.6691
17.7573 | -35.0283 | -27.6766
17.7634 | -35.0287 | -27.6838
17.7695 | -35.0293 | -27.6909
17.7756 | -35.0297 | -27.6983
17.7817 | -35.0302 | -27.7058
17.7878 | -35.0307 | -27.7131
17.7939 | -35.0312 | -27.7205
17.8 | -35.0317 | -27.7278
17.8061 | -35.0322 | -27.7351
17.8122 | -35.0327 | -27.7424
17.8183 | -35.0332 | -27.7497
17.8244 | -35.0337 | -27.757
17.8305 | -35.0342 | -27.7644
17.8366 | -35.0347 | -27.7715
17.8428 | -35.0352 | -27.7788
17.8489 | -35.0357 | -27.7861
17.855 | -35.0362 | -27.7934
17.8611 | -35.0367 | -27.8007
17.8672 | -35.0372 | -27.808
17.8733 | -35.0377 | -27.8153
17.8794 | -35.0382 | -27.8226
17.8855 | -35.0387 | -27.83
17.8916 | -35.0392 | -27.8374
17.8977 | -35.0397 | -27.8447
17.9038 | -35.0402 | -27.8518
17.9099 | -35.0407 | -27.8592
17.916 | -35.0412 | -27.8666
17.9221 | -35.0416 | -27.8741
17.9282 | -35.0421 | -27.8813
17.9343 | -35.0426 | -27.8886
17.9404 | -35.0431 | -27.8957
17.9465 | -35.0436 | -27.9032
17.9526 | -35.0441 | -27.9106
17.9587 | -35.0446 | -27.9181
17.9648 | -35.0451 | -27.9254
17.9709 | -35.0456 | -27.9327
17.977 | -35.0461 | -27.9401
17.9832 | -35.0466 | -27.9474
17.9893 | -35.0471 | -27.9547
17.9954 | -35.0475 | -27.9621
18.0015 | -35.048 | -27.9694
18.0076 | -35.0485 | -27.9765
18.0137 | -35.049 | -27.9837
18.0198 | -35.0495 | -27.9911
18.0259 | -35.05 | -27.9985
18.032 | -35.0505 | -28.0059
18.0381 | -35.051 | -28.0133
18.0442 | -35.0515 | -28.0206
18.0503 | -35.052 | -28.0278
18.0564 | -35.0525 | -28.0353
18.0625 | -35.0529 | -28.0426
18.0686 | -35.0534 | -28.0499
18.0747 | -35.0539 | -28.0573
18.0808 | -35.0544 | -28.0647
18.0869 | -35.0549 | -28.0719
18.093 | -35.0554 | -28.0793
18.0991 | -35.0558 | -28.0866
18.1052 | -35.0563 | -28.094
18.1113 | -35.0568 | -28.1012
18.1174 | -35.0573 | -28.1083
18.1236 | -35.0578 | -28.1157
18.1297 | -35.0583 | -28.1229
18.1358 | -35.0588 | -28.13
18.1419 | -35.0593 | -28.1374
18.148 | -35.0598 | -28.1445
18.1541 | -35.0603 | -28.1517
18.1602 | -35.0607 | -28.1589
18.1663 | -35.0612 | -28.1659
18.1724 | -35.0617 | -28.1732
18.1785 | -35.0622 | -28.1803
18.1846 | -35.0627 | -28.1874
18.1907 | -35.0632 | -28.1947
18.1968 | -35.0637 | -28.2021
18.2029 | -35.0642 | -28.2094
18.209 | -35.0646 | -28.2169
18.2151 | -35.0651 | -28.2242
18.2212 | -35.0656 | -28.2314
18.2273 | -35.0661 | -28.2388
18.2334 | -35.0666 | -28.2462
18.2395 | -35.0671 | -28.2535
18.2456 | -35.0676 | -28.2608
18.2517 | -35.0681 | -28.2681
18.2578 | -35.0685 | -28.2754
18.2639 | -35.069 | -28.2827
18.2701 | -35.0695 | -28.29
18.2762 | -35.07 | -28.2973
18.2823 | -35.0705 | -28.3046
18.2884 | -35.071 | -28.3118
18.2945 | -35.0715 | -28.3191
18.3006 | -35.072 | -28.3265
18.3067 | -35.0725 | -28.3338
18.3128 | -35.073 | -28.3412
18.3189 | -35.0734 | -28.3486
18.325 | -35.0739 | -28.3556
18.3311 | -35.0744 | -28.3628
18.3372 | -35.0749 | -28.3702
18.3433 | -35.0753 | -28.3776
18.3494 | -35.0758 | -28.3848
18.3555 | -35.0763 | -28.3921
18.3616 | -35.0768 | -28.3992
18.3677 | -35.0773 | -28.4064
18.3738 | -35.0778 | -28.4134
18.3799 | -35.0783 | -28.4206
18.386 | -35.0788 | -28.428
18.3921 | -35.0792 | -28.4353
18.3982 | -35.0797 | -28.4423
18.4043 | -35.0802 | -28.4494
18.4105 | -35.0806 | -28.4567
18.4166 | -35.0811 | -28.4638
18.4227 | -35.0816 | -28.471
18.4288 | -35.0821 | -28.4783
18.4349 | -35.0826 | -28.4853
18.441 | -35.083 | -28.4926
18.4471 | -35.0835 | -28.4998
18.4532 | -35.084 | -28.5072
18.4593 | -35.0845 | -28.5144
18.4654 | -35.0849 | -28.5214
18.4715 | -35.0854 | -28.5286
18.4776 | -35.0859 | -28.5359
18.4837 | -35.0864 | -28.5431
18.4898 | -35.0868 | -28.5502
18.4959 | -35.0873 | -28.5577
18.502 | -35.0878 | -28.565
18.5081 | -35.0882 | -28.5722
18.5142 | -35.0887 | -28.5795
18.5203 | -35.0892 | -28.5867
18.5264 | -35.0897 | -28.5939
18.5325 | -35.0902 | -28.6011
18.5386 | -35.0907 | -28.6083
18.5447 | -35.0912 | -28.6156
18.5508 | -35.0917 | -28.6228
18.557 | -35.0922 | -28.6301
18.5631 | -35.0927 | -28.6374
18.5692 | -35.0931 | -28.6446
18.5753 | -35.0936 | -28.6519
18.5814 | -35.0941 | -28.659
18.5875 | -35.0946 | -28.6662
18.5936 | -35.0951 | -28.6734
18.5997 | -35.0955 | -28.6807
18.6058 | -35.096 | -28.6879
18.6119 | -35.0965 | -28.6951
18.618 | -35.097 | -28.7023
18.6241 | -35.0975 | -28.7095
18.6302 | -35.098 | -28.7168
18.6363 | -35.0984 | -28.724
18.6424 | -35.0989 | -28.7311
18.6485 | -35.0994 | -28.7383
18.6546 | -35.0998 | -28.7455
18.6607 | -35.1003 | -28.7528
18.6668 | -35.1008 | -28.7601
18.6729 | -35.1013 | -28.7673
18.679 | -35.1017 | -28.7745
18.6851 | -35.1022 | -28.7818
18.6912 | -35.1027 | -28.789
18.6974 | -35.1031 | -28.7961
18.7035 | -35.1036 | -28.8033
18.7096 | -35.1041 | -28.8105
18.7157 | -35.1046 | -28.8176
18.7218 | -35.1051 | -28.8249
18.7279 | -35.1056 | -28.8322
18.734 | -35.1061 | -28.8394
18.7401 | -35.1066 | -28.8466
18.7462 | -35.107 | -28.8533
18.7523 | -35.1075 | -28.8606
18.7584 | -35.108 | -28.8679
18.7645 | -35.1085 | -28.875
18.7706 | -35.1089 | -28.8821
18.7767 | -35.1094 | -28.8892
18.7828 | -35.1099 | -28.8961
18.7889 | -35.1104 | -28.9033
18.795 | -35.1109 | -28.9105
18.8011 | -35.1113 | -28.9176
18.8072 | -35.1118 | -28.9248
18.8133 | -35.1123 | -28.9321
18.8194 | -35.1127 | -28.939
18.8255 | -35.1132 | -28.9461
18.8316 | -35.1137 | -28.9533
18.8377 | -35.1142 | -28.9605
18.8439 | -35.1147 | -28.9677
18.85 | -35.1151 | -28.9749
18.8561 | -35.1156 | -28.9819
18.8622 | -35.1161 | -28.9892
18.8683 | -35.1166 | -28.9963
18.8744 | -35.117 | -29.0034
18.8805 | -35.1175 | -29.0105
18.8866 | -35.118 | -29.0177
18.8927 | -35.1185 | -29.0248
18.8988 | -35.119 | -29.0319
18.9049 | -35.1194 | -29.0391
18.911 | -35.1199 | -29.0462
18.9171 | -35.1204 | -29.0534
18.9232 | -35.1209 | -29.0606
18.9293 | -35.1214 | -29.0678
18.9354 | -35.1219 | -29.075
18.9415 | -35.1224 | -29.0821
18.9476 | -35.1228 | -29.0891
18.9537 | -35.1233 | -29.0962
18.9598 | -35.1237 | -29.1031
18.9659 | -35.1242 | -29.1102
18.972 | -35.1247 | -29.1174
18.9781 | -35.1251 | -29.1245
18.9843 | -35.1256 | -29.1319
18.9904 | -35.1261 | -29.1389
18.9965 | -35.1266 | -29.1458
19.0026 | -35.127 | -29.153
19.0087 | -35.1275 | -29.1602
19.0148 | -35.128 | -29.1674
19.0209 | -35.1285 | -29.1746
19.027 | -35.129 | -29.1818
19.0331 | -35.1295 | -29.1887
19.0392 | -35.13 | -29.196
19.0453 | -35.1304 | -29.203
19.0514 | -35.1309 | -29.2101
19.0575 | -35.1314 | -29.2174
19.0636 | -35.1319 | -29.2243
19.0697 | -35.1324 | -29.2313
19.0758 | -35.1329 | -29.2384
19.0819 | -35.1333 | -29.2456
19.088 | -35.1338 | -29.2527
19.0941 | -35.1343 | -29.2599
19.1002 | -35.1348 | -29.2668
19.1063 | -35.1352 | -29.2738
19.1124 | -35.1357 | -29.2809
19.1185 | -35.1362 | -29.2879
19.1246 | -35.1366 | -29.2949
19.1308 | -35.1371 | -29.302
19.1369 | -35.1376 | -29.3091
19.143 | -35.138 | -29.3162
19.1491 | -35.1385 | -29.3234
19.1552 | -35.139 | -29.3306
19.1613 | -35.1395 | -29.3378
19.1674 | -35.1399 | -29.3448
19.1735 | -35.1404 | -29.3517
19.1796 | -35.1409 | -29.3589
19.1857 | -35.1413 | -29.3659
19.1918 | -35.1418 | -29.373
19.1979 | -35.1423 | -29.38
19.204 | -35.1428 | -29.3868
19.2101 | -35.1433 | -29.3939
19.2162 | -35.1438 | -29.4008
19.2223 | -35.1442 | -29.408
19.2284 | -35.1447 | -29.4152
19.2345 | -35.1451 | -29.4225
19.2406 | -35.1456 | -29.4297
19.2467 | -35.1461 | -29.4368
19.2528 | -35.1465 | -29.444
19.2589 | -35.147 | -29.4511
19.265 | -35.1475 | -29.4584
19.2712 | -35.1479 | -29.4654
19.2773 | -35.1484 | -29.4723
19.2834 | -35.1489 | -29.4791
19.2895 | -35.1493 | -29.4861
19.2956 | -35.1498 | -29.4932
19.3017 | -35.1503 | -29.5002
19.3078 | -35.1507 | -29.5073
19.3139 | -35.1512 | -29.5141
19.32 | -35.1517 | -29.5212
19.3261 | -35.1521 | -29.5282
19.3322 | -35.1526 | -29.5351
19.3383 | -35.1531 | -29.5422
19.3444 | -35.1535 | -29.5489
19.3505 | -35.154 | -29.5559
19.3566 | -35.1545 | -29.5632
19.3627 | -35.1549 | -29.5701
19.3688 | -35.1554 | -29.5772
19.3749 | -35.1559 | -29.5842
19.381 | -35.1563 | -29.5909
19.3871 | -35.1568 | -29.5979
19.3932 | -35.1573 | -29.605
19.3993 | -35.1577 | -29.6121
19.4054 | -35.1582 | -29.6193
19.4115 | -35.1587 | -29.6263
19.4177 | -35.1591 | -29.6332
19.4238 | -35.1596 | -29.6402
19.4299 | -35.1601 | -29.6473
19.436 | -35.1606 | -29.6544
19.4421 | -35.161 | -29.6613
19.4482 | -35.1615 | -29.6684
19.4543 | -35.162 | -29.6752
19.4604 | -35.1625 | -29.6824
19.4665 | -35.1629 | -29.6896
19.4726 | -35.1634 | -29.6965
19.4787 | -35.1639 | -29.7038
19.4848 | -35.1644 | -29.7107
19.4909 | -35.1649 | -29.7176
19.497 | -35.1653 | -29.7245
19.5031 | -35.1658 | -29.7314
19.5092 | -35.1663 | -29.7385
19.5153 | -35.1668 | -29.7454
19.5214 | -35.1673 | -29.7521
19.5275 | -35.1677 | -29.7591
19.5336 | -35.1682 | -29.766
19.5397 | -35.1687 | -29.7729
19.5458 | -35.1691 | -29.7799
19.5519 | -35.1696 | -29.787
19.5581 | -35.1701 | -29.7938
19.5642 | -35.1705 | -29.8009
19.5703 | -35.1709 | -29.8082
19.5764 | -35.1714 | -29.8153
19.5825 | -35.1719 | -29.8221
19.5886 | -35.1724 | -29.8292
19.5947 | -35.1729 | -29.8361
19.6008 | -35.1733 | -29.8429
19.6069 | -35.1738 | -29.8499
19.613 | -35.1742 | -29.8569
19.6191 | -35.1747 | -29.8639
19.6252 | -35.1752 | -29.8708
19.6313 | -35.1757 | -29.8777
19.6374 | -35.1761 | -29.8848
19.6435 | -35.1766 | -29.8917
19.6496 | -35.1771 | -29.8987
19.6557 | -35.1776 | -29.9057
19.6618 | -35.178 | -29.9125
19.6679 | -35.1785 | -29.9195
19.674 | -35.179 | -29.9264
19.6801 | -35.1795 | -29.9334
19.6862 | -35.18 | -29.9401
19.6923 | -35.1804 | -29.9469
19.6984 | -35.1809 | -29.9535
19.7046 | -35.1814 | -29.9604
19.7107 | -35.1819 | -29.9673
19.7168 | -35.1824 | -29.9743
19.7229 | -35.1829 | -29.9814
19.729 | -35.1834 | -29.9884
19.7351 | -35.1838 | -29.9952
19.7412 | -35.1843 | -30.0021
19.7473 | -35.1848 | -30.0092
19.7534 | -35.1853 | -30.0161
19.7595 | -35.1857 | -30.023
19.7656 | -35.1862 | -30.0299
19.7717 | -35.1867 | -30.0367
19.7778 | -35.1871 | -30.0437
19.7839 | -35.1876 | -30.0507
19.79 | -35.1881 | -30.0575
19.7961 | -35.1885 | -30.0644
19.8022 | -35.189 | -30.0711
19.8083 | -35.1895 | -30.0779
19.8144 | -35.19 | -30.0847
19.8205 | -35.1904 | -30.0917
19.8266 | -35.1909 | -30.0988
19.8327 | -35.1913 | -30.1059
19.8388 | -35.1918 | -30.1129
19.845 | -35.1922 | -30.1198
19.8511 | -35.1927 | -30.1267
19.8572 | -35.1932 | -30.1336
19.8633 | -35.1936 | -30.1407
19.8694 | -35.1941 | -30.1477
19.8755 | -35.1945 | -30.1547
19.8816 | -35.1949 | -30.1618
19.8877 | -35.1954 | -30.169
19.8938 | -35.1959 | -30.1759
19.8999 | -35.1964 | -30.1829
19.906 | -35.1968 | -30.19
19.9121 | -35.1973 | -30.1967
19.9182 | -35.1978 | -30.2034
19.9243 | -35.1983 | -30.2101
19.9304 | -35.1987 | -30.217
19.9365 | -35.1992 | -30.224
19.9426 | -35.1997 | -30.2308
19.9487 | -35.2002 | -30.2379
19.9548 | -35.2006 | -30.2448
19.9609 | -35.2011 | -30.2517
19.967 | -35.2016 | -30.2586
19.9731 | -35.2021 | -30.2655
19.9792 | -35.2025 | -30.2725
19.9853 | -35.203 | -30.2795
19.9915 | -35.2035 | -30.2866
19.9976 | -35.204 | -30.2935
20.0037 | -35.2044 | -30.3005
20.0098 | -35.2049 | -30.3075
20.0159 | -35.2054 | -30.3142
20.022 | -35.2058 | -30.321
20.0281 | -35.2063 | -30.3278
20.0342 | -35.2067 | -30.3349
20.0403 | -35.2072 | -30.3417
20.0464 | -35.2077 | -30.3486
20.0525 | -35.2082 | -30.3552
20.0586 | -35.2086 | -30.3621
20.0647 | -35.2091 | -30.369
20.0708 | -35.2096 | -30.376
20.0769 | -35.21 | -30.383
20.083 | -35.2104 | -30.3896
20.0891 | -35.2109 | -30.3964
20.0952 | -35.2114 | -30.4033
20.1013 | -35.2118 | -30.4104
20.1074 | -35.2123 | -30.4172
20.1135 | -35.2128 | -30.4239
20.1196 | -35.2132 | -30.4304
20.1257 | -35.2137 | -30.4373
20.1319 | -35.2141 | -30.444
20.138 | -35.2146 | -30.4509
20.1441 | -35.215 | -30.4576
20.1502 | -35.2155 | -30.4646
20.1563 | -35.216 | -30.4711
20.1624 | -35.2165 | -30.4779
20.1685 | -35.2169 | -30.4848
20.1746 | -35.2174 | -30.4916
20.1807 | -35.2179 | -30.4985
20.1868 | -35.2184 | -30.5054
20.1929 | -35.2188 | -30.5122
20.199 | -35.2192 | -30.519
20.2051 | -35.2196 | -30.5257
20.2112 | -35.2201 | -30.5325
20.2173 | -35.2206 | -30.5393
20.2234 | -35.221 | -30.5463
20.2295 | -35.2215 | -30.5528
20.2356 | -35.222 | -30.5596
20.2417 | -35.2224 | -30.5664
20.2478 | -35.2229 | -30.5731
20.2539 | -35.2233 | -30.58
20.26 | -35.2238 | -30.5869
20.2661 | -35.2243 | -30.5936
20.2723 | -35.2248 | -30.6004
20.2784 | -35.2253 | -30.6072
20.2845 | -35.2257 | -30.614
20.2906 | -35.2262 | -30.6208
20.2967 | -35.2267 | -30.6276
20.3028 | -35.2271 | -30.6343
20.3089 | -35.2276 | -30.641
20.315 | -35.2281 | -30.6479
20.3211 | -35.2286 | -30.6547
20.3272 | -35.229 | -30.6616
20.3333 | -35.2295 | -30.6682
20.3394 | -35.23 | -30.6752
20.3455 | -35.2304 | -30.6822
20.3516 | -35.2309 | -30.6888
20.3577 | -35.2313 | -30.6955
20.3638 | -35.2318 | -30.7024
20.3699 | -35.2322 | -30.7089
20.376 | -35.2327 | -30.7156
20.3821 | -35.2332 | -30.7227
20.3882 | -35.2337 | -30.7295
20.3943 | -35.2341 | -30.7363
20.4004 | -35.2346 | -30.7427
20.4065 | -35.235 | -30.7492
20.4126 | -35.2355 | -30.7562
20.4188 | -35.236 | -30.763
20.4249 | -35.2364 | -30.7696
20.431 | -35.2369 | -30.7763
20.4371 | -35.2373 | -30.7827
20.4432 | -35.2378 | -30.7897
20.4493 | -35.2382 | -30.7965
20.4554 | -35.2387 | -30.8031
20.4615 | -35.2392 | -30.81
20.4676 | -35.2396 | -30.8169
20.4737 | -35.2402 | -30.8234
20.4798 | -35.2406 | -30.8303
20.4859 | -35.2411 | -30.8371
20.492 | -35.2416 | -30.8436
20.4981 | -35.242 | -30.8503
20.5042 | -35.2425 | -30.857
20.5103 | -35.2429 | -30.8634
20.5164 | -35.2434 | -30.8701
20.5225 | -35.2438 | -30.877
20.5286 | -35.2443 | -30.8838
20.5347 | -35.2448 | -30.8905
20.5408 | -35.2452 | -30.8973
20.5469 | -35.2456 | -30.904
20.553 | -35.2461 | -30.9108
20.5592 | -35.2466 | -30.9176
20.5653 | -35.2471 | -30.9245
20.5714 | -35.2475 | -30.9312
20.5775 | -35.248 | -30.9377
20.5836 | -35.2484 | -30.9445
20.5897 | -35.2489 | -30.9513
20.5958 | -35.2494 | -30.9581
20.6019 | -35.2498 | -30.9649
20.608 | -35.2503 | -30.9717
20.6141 | -35.2508 | -30.9779
20.6202 | -35.2512 | -30.9847
20.6263 | -35.2517 | -30.9913
20.6324 | -35.2522 | -30.9981
20.6385 | -35.2526 | -31.0048
20.6446 | -35.2531 | -31.0115
20.6507 | -35.2535 | -31.0179
20.6568 | -35.254 | -31.0246
20.6629 | -35.2544 | -31.0313
20.669 | -35.2549 | -31.0379
20.6751 | -35.2554 | -31.0445
20.6812 | -35.2559 | -31.051
20.6873 | -35.2563 | -31.0577
20.6934 | -35.2568 | -31.0645
20.6995 | -35.2573 | -31.0712
20.7057 | -35.2577 | -31.0781
20.7118 | -35.2582 | -31.0847
20.7179 | -35.2586 | -31.0911
20.724 | -35.2591 | -31.0978
20.7301 | -35.2595 | -31.1044
20.7362 | -35.26 | -31.111
20.7423 | -35.2605 | -31.1176
20.7484 | -35.2609 | -31.1245
20.7545 | -35.2614 | -31.1307
20.7606 | -35.2619 | -31.1375
20.7667 | -35.2624 | -31.1441
20.7728 | -35.2628 | -31.1508
20.7789 | -35.2633 | -31.1575
20.785 | -35.2637 | -31.1642
20.7911 | -35.2642 | -31.171
20.7972 | -35.2646 | -31.1777
20.8033 | -35.2651 | -31.1842
20.8094 | -35.2655 | -31.1909
20.8155 | -35.266 | -31.1977
20.8216 | -35.2664 | -31.2044
20.8277 | -35.2669 | -31.2108
20.8338 | -35.2673 | -31.2173
20.8399 | -35.2678 | -31.224
20.8461 | -35.2682 | -31.2308
20.8522 | -35.2687 | -31.2376
20.8583 | -35.2692 | -31.244
20.8644 | -35.2696 | -31.2508
20.8705 | -35.2701 | -31.2574
20.8766 | -35.2705 | -31.2642
20.8827 | -35.271 | -31.2709
20.8888 | -35.2714 | -31.2773
20.8949 | -35.2719 | -31.2838
20.901 | -35.2723 | -31.2906
20.9071 | -35.2728 | -31.297
20.9132 | -35.2732 | -31.3037
20.9193 | -35.2737 | -31.3105
20.9254 | -35.2742 | -31.3172
20.9315 | -35.2747 | -31.3235
20.9376 | -35.2751 | -31.3302
20.9437 | -35.2756 | -31.3368
20.9498 | -35.276 | -31.3432
20.9559 | -35.2765 | -31.3496
20.962 | -35.277 | -31.3564
20.9681 | -35.2774 | -31.3629
20.9742 | -35.2779 | -31.3696
20.9803 | -35.2784 | -31.3763
20.9864 | -35.2788 | -31.3832
20.9926 | -35.2793 | -31.39
20.9987 | -35.2797 | -31.3965
21.0048 | -35.2802 | -31.4031
21.0109 | -35.2807 | -31.4099
21.017 | -35.2812 | -31.4164
21.0231 | -35.2816 | -31.4232
21.0292 | -35.2821 | -31.4296
21.0353 | -35.2826 | -31.4361
21.0414 | -35.2831 | -31.4426
21.0475 | -35.2835 | -31.4492
21.0536 | -35.284 | -31.4558
21.0597 | -35.2844 | -31.4624
21.0658 | -35.2849 | -31.4689
21.0719 | -35.2853 | -31.4758
21.078 | -35.2858 | -31.4826
21.0841 | -35.2862 | -31.4894
21.0902 | -35.2866 | -31.4962
21.0963 | -35.2871 | -31.503
21.1024 | -35.2876 | -31.5097
21.1085 | -35.288 | -31.5161
21.1146 | -35.2885 | -31.5228
21.1207 | -35.2889 | -31.5295
21.1268 | -35.2894 | -31.536
21.133 | -35.2899 | -31.5426
21.1391 | -35.2903 | -31.5493
21.1452 | -35.2908 | -31.5558
21.1513 | -35.2912 | -31.5625
21.1574 | -35.2917 | -31.569
21.1635 | -35.2922 | -31.5756
21.1696 | -35.2926 | -31.5822
21.1757 | -35.293 | -31.5886
21.1818 | -35.2935 | -31.5952
21.1879 | -35.2939 | -31.6018
21.194 | -35.2943 | -31.6084
21.2001 | -35.2948 | -31.615
21.2062 | -35.2952 | -31.6218
21.2123 | -35.2957 | -31.6284
21.2184 | -35.2962 | -31.6349
21.2245 | -35.2966 | -31.6415
21.2306 | -35.297 | -31.6479
21.2367 | -35.2975 | -31.6548
21.2428 | -35.298 | -31.6614
21.2489 | -35.2984 | -31.6677
21.255 | -35.2989 | -31.6745
21.2611 | -35.2993 | -31.6811
21.2672 | -35.2998 | -31.6877
21.2733 | -35.3002 | -31.6944
21.2795 | -35.3006 | -31.7006
21.2856 | -35.3011 | -31.7071
21.2917 | -35.3016 | -31.7136
21.2978 | -35.3021 | -31.7201
21.3039 | -35.3025 | -31.7267
21.31 | -35.303 | -31.7333
21.3161 | -35.3034 | -31.7394
21.3222 | -35.3039 | -31.746
21.3283 | -35.3043 | -31.7526
21.3344 | -35.3048 | -31.7593
21.3405 | -35.3053 | -31.7659
21.3466 | -35.3058 | -31.7724
21.3527 | -35.3062 | -31.779
21.3588 | -35.3067 | -31.7855
21.3649 | -35.3071 | -31.792
21.371 | -35.3076 | -31.7983
21.3771 | -35.308 | -31.8049
21.3832 | -35.3085 | -31.8113
21.3893 | -35.3089 | -31.8178
21.3954 | -35.3093 | -31.8245
21.4015 | -35.3098 | -31.831
21.4076 | -35.3103 | -31.8375
21.4137 | -35.3108 | -31.844
21.4199 | -35.3112 | -31.8507
21.426 | -35.3116 | -31.8572
21.4321 | -35.3121 | -31.8637
21.4382 | -35.3125 | -31.8703
21.4443 | -35.313 | -31.8769
21.4504 | -35.3134 | -31.8836
21.4565 | -35.3139 | -31.8901
21.4626 | -35.3143 | -31.8968
21.4687 | -35.3148 | -31.9035
21.4748 | -35.3153 | -31.9099
21.4809 | -35.3157 | -31.9166
21.487 | -35.3162 | -31.923
21.4931 | -35.3166 | -31.9292
21.4992 | -35.3171 | -31.9357
21.5053 | -35.3175 | -31.9421
21.5114 | -35.318 | -31.9489
21.5175 | -35.3184 | -31.9554
21.5236 | -35.3188 | -31.962
21.5297 | -35.3193 | -31.9685
21.5358 | -35.3197 | -31.975
21.5419 | -35.3201 | -31.9815
21.548 | -35.3206 | -31.9882
21.5541 | -35.3211 | -31.9949
21.5602 | -35.3215 | -32.0018
21.5664 | -35.322 | -32.008
21.5725 | -35.3225 | -32.0145
21.5786 | -35.3229 | -32.0212
21.5847 | -35.3233 | -32.0278
21.5908 | -35.3238 | -32.0343
21.5969 | -35.3242 | -32.0406
21.603 | -35.3247 | -32.0472
21.6091 | -35.3251 | -32.0537
21.6152 | -35.3256 | -32.0602
21.6213 | -35.326 | -32.0668
21.6274 | -35.3265 | -32.0733
21.6335 | -35.327 | -32.0797
21.6396 | -35.3274 | -32.0863
21.6457 | -35.3279 | -32.0928
21.6518 | -35.3283 | -32.0993
21.6579 | -35.3288 | -32.106
21.664 | -35.3292 | -32.1125
21.6701 | -35.3297 | -32.1187
21.6762 | -35.3302 | -32.1253
21.6823 | -35.3306 | -32.1322
21.6884 | -35.331 | -32.1387
21.6945 | -35.3314 | -32.1451
21.7006 | -35.3319 | -32.1517
21.7068 | -35.3323 | -32.1579
21.7129 | -35.3328 | -32.1644
21.719 | -35.3332 | -32.1711
21.7251 | -35.3336 | -32.1776
21.7312 | -35.3341 | -32.1842
21.7373 | -35.3345 | -32.1906
21.7434 | -35.335 | -32.1972
21.7495 | -35.3354 | -32.2037
21.7556 | -35.3359 | -32.2102
21.7617 | -35.3364 | -32.2168
21.7678 | -35.3368 | -32.2231
21.7739 | -35.3373 | -32.2294
21.78 | -35.3377 | -32.2359
21.7861 | -35.3382 | -32.2424
21.7922 | -35.3386 | -32.2491
21.7983 | -35.339 | -32.2556
21.8044 | -35.3395 | -32.262
21.8105 | -35.34 | -32.2686
21.8166 | -35.3404 | -32.2751
21.8227 | -35.3409 | -32.2818
21.8288 | -35.3413 | -32.2884
21.8349 | -35.3418 | -32.295
21.841 | -35.3422 | -32.3017
21.8471 | -35.3426 | -32.3077
21.8533 | -35.3431 | -32.314
21.8594 | -35.3436 | -32.3206
21.8655 | -35.344 | -32.3271
21.8716 | -35.3444 | -32.3339
21.8777 | -35.3448 | -32.3404
21.8838 | -35.3453 | -32.3469
21.8899 | -35.3458 | -32.3537
21.896 | -35.3462 | -32.3601
21.9021 | -35.3467 | -32.3667
21.9082 | -35.3471 | -32.3733
21.9143 | -35.3475 | -32.3796
21.9204 | -35.348 | -32.3861
21.9265 | -35.3484 | -32.3926
21.9326 | -35.3489 | -32.3993
21.9387 | -35.3493 | -32.4059
21.9448 | -35.3497 | -32.4122
21.9509 | -35.3502 | -32.4186
21.957 | -35.3506 | -32.4252
21.9631 | -35.3511 | -32.4317
21.9692 | -35.3515 | -32.4379
21.9753 | -35.3519 | -32.4442
21.9814 | -35.3523 | -32.4506
21.9875 | -35.3528 | -32.4568
21.9937 | -35.3532 | -32.4632
21.9998 | -35.3537 | -32.4698
22.0059 | -35.3541 | -32.4764
22.012 | -35.3546 | -32.4828
22.0181 | -35.355 | -32.4892
22.0242 | -35.3554 | -32.4957
22.0303 | -35.3559 | -32.5022
22.0364 | -35.3562 | -32.5086
22.0425 | -35.3567 | -32.5149
22.0486 | -35.3571 | -32.5214
22.0547 | -35.3575 | -32.5279
22.0608 | -35.358 | -32.5343
22.0669 | -35.3584 | -32.5409
22.073 | -35.3589 | -32.5473
22.0791 | -35.3593 | -32.5539
22.0852 | -35.3597 | -32.5606
22.0913 | -35.3601 | -32.5669
22.0974 | -35.3606 | -32.5734
22.1035 | -35.361 | -32.58
22.1096 | -35.3614 | -32.5867
22.1157 | -35.3619 | -32.5933
22.1218 | -35.3623 | -32.5998
22.1279 | -35.3627 | -32.6061
22.134 | -35.3632 | -32.6128
22.1402 | -35.3636 | -32.6193
22.1463 | -35.3641 | -32.6259
22.1524 | -35.3645 | -32.6322
22.1585 | -35.3649 | -32.6387
22.1646 | -35.3654 | -32.645
22.1707 | -35.3658 | -32.6516
22.1768 | -35.3662 | -32.6581
22.1829 | -35.3667 | -32.6645
22.189 | -35.3671 | -32.6709
22.1951 | -35.3675 | -32.6773
22.2012 | -35.368 | -32.6838
22.2073 | -35.3684 | -32.6903
22.2134 | -35.3688 | -32.6967
22.2195 | -35.3693 | -32.7031
22.2256 | -35.3697 | -32.7094
22.2317 | -35.3702 | -32.7156
22.2378 | -35.3706 | -32.7223
22.2439 | -35.371 | -32.7287
22.25 | -35.3714 | -32.7351
22.2561 | -35.3719 | -32.7413
22.2622 | -35.3723 | -32.7477
22.2683 | -35.3727 | -32.754
22.2744 | -35.3732 | -32.7604
22.2806 | -35.3736 | -32.7669
22.2867 | -35.374 | -32.7734
22.2928 | -35.3745 | -32.7799
22.2989 | -35.3749 | -32.7864
22.305 | -35.3754 | -32.7927
22.3111 | -35.3758 | -32.7993
22.3172 | -35.3763 | -32.8059
22.3233 | -35.3767 | -32.8122
22.3294 | -35.3772 | -32.8188
22.3355 | -35.3776 | -32.8251
22.3416 | -35.378 | -32.8315
22.3477 | -35.3784 | -32.8377
22.3538 | -35.3788 | -32.8444
22.3599 | -35.3793 | -32.8508
22.366 | -35.3797 | -32.8572
22.3721 | -35.3802 | -32.8635
22.3782 | -35.3806 | -32.8699
22.3843 | -35.381 | -32.8762
22.3904 | -35.3815 | -32.8829
22.3965 | -35.3819 | -32.8893
22.4026 | -35.3823 | -32.8957
22.4087 | -35.3828 | -32.902
22.4148 | -35.3832 | -32.9085
22.4209 | -35.3836 | -32.9149
22.4271 | -35.3841 | -32.9214
22.4332 | -35.3845 | -32.9278
22.4393 | -35.3849 | -32.9341
22.4454 | -35.3854 | -32.9402
22.4515 | -35.3858 | -32.9466
22.4576 | -35.3862 | -32.953
22.4637 | -35.3867 | -32.9596
22.4698 | -35.3871 | -32.966
22.4759 | -35.3875 | -32.9723
22.482 | -35.388 | -32.9789
22.4881 | -35.3884 | -32.9854
22.4942 | -35.3888 | -32.9919
22.5003 | -35.3893 | -32.9984
22.5064 | -35.3897 | -33.0048
22.5125 | -35.3901 | -33.0113
22.5186 | -35.3905 | -33.0177
22.5247 | -35.3909 | -33.024
22.5308 | -35.3914 | -33.0305
22.5369 | -35.3918 | -33.0369
22.543 | -35.3922 | -33.0435
22.5491 | -35.3926 | -33.0498
22.5552 | -35.3931 | -33.0563
22.5613 | -35.3935 | -33.0627
22.5675 | -35.3939 | -33.0691
22.5736 | -35.3944 | -33.0757
22.5797 | -35.3948 | -33.082
22.5858 | -35.3952 | -33.0884
22.5919 | -35.3957 | -33.0947
22.598 | -35.3961 | -33.1012
22.6041 | -35.3965 | -33.1076
22.6102 | -35.3969 | -33.1138
22.6163 | -35.3974 | -33.12
22.6224 | -35.3978 | -33.1264
22.6285 | -35.3982 | -33.1328
22.6346 | -35.3986 | -33.1392
22.6407 | -35.3991 | -33.1455
22.6468 | -35.3995 | -33.1519
22.6529 | -35.3999 | -33.1583
22.659 | -35.4003 | -33.1646
22.6651 | -35.4007 | -33.1711
22.6712 | -35.4012 | -33.1774
22.6773 | -35.4016 | -33.1839
22.6834 | -35.402 | -33.1903
22.6895 | -35.4025 | -33.1965
22.6956 | -35.4029 | -33.2029
22.7017 | -35.4033 | -33.2093
22.7079 | -35.4037 | -33.2156
22.714 | -35.4042 | -33.2217
22.7201 | -35.4046 | -33.2281
22.7262 | -35.405 | -33.2345
22.7323 | -35.4054 | -33.2409
22.7384 | -35.4059 | -33.2472
22.7445 | -35.4063 | -33.2536
22.7506 | -35.4067 | -33.2599
22.7567 | -35.4071 | -33.2661
22.7628 | -35.4075 | -33.2724
22.7689 | -35.4079 | -33.2788
22.775 | -35.4084 | -33.2851
22.7811 | -35.4088 | -33.2916
22.7872 | -35.4092 | -33.2979
22.7933 | -35.4097 | -33.304
22.7994 | -35.4101 | -33.3101
22.8055 | -35.4105 | -33.3164
22.8116 | -35.411 | -33.3227
22.8177 | -35.4114 | -33.3293
22.8238 | -35.4118 | -33.3357
22.8299 | -35.4122 | -33.3419
22.836 | -35.4127 | -33.3481
22.8421 | -35.4131 | -33.3545
22.8482 | -35.4135 | -33.3608
22.8544 | -35.4139 | -33.3671
22.8605 | -35.4143 | -33.3735
22.8666 | -35.4147 | -33.3795
22.8727 | -35.4152 | -33.3858
22.8788 | -35.4156 | -33.3922
22.8849 | -35.416 | -33.3985
22.891 | -35.4165 | -33.4048
22.8971 | -35.4169 | -33.411
22.9032 | -35.4173 | -33.4171
22.9093 | -35.4178 | -33.4235
22.9154 | -35.4182 | -33.4298
22.9215 | -35.4186 | -33.436
22.9276 | -35.419 | -33.4424
22.9337 | -35.4194 | -33.4485
22.9398 | -35.4199 | -33.4548
22.9459 | -35.4203 | -33.4611
22.952 | -35.4207 | -33.4675
22.9581 | -35.4212 | -33.4738
22.9642 | -35.4216 | -33.4802
22.9703 | -35.422 | -33.4863
22.9764 | -35.4224 | -33.4926
22.9825 | -35.4228 | -33.4988
22.9886 | -35.4233 | -33.5051
22.9948 | -35.4237 | -33.5114
23.0009 | -35.4241 | -33.5178
23.007 | -35.4246 | -33.5239
23.0131 | -35.425 | -33.5301
23.0192 | -35.4254 | -33.5365
23.0253 | -35.4258 | -33.543
23.0314 | -35.4262 | -33.5493
23.0375 | -35.4266 | -33.5557
23.0436 | -35.427 | -33.5619
23.0497 | -35.4275 | -33.5684
23.0558 | -35.4278 | -33.5748
23.0619 | -35.4283 | -33.5812
23.068 | -35.4287 | -33.5874
23.0741 | -35.4291 | -33.5935
23.0802 | -35.4295 | -33.5998
23.0863 | -35.4299 | -33.6062
23.0924 | -35.4304 | -33.6126
23.0985 | -35.4308 | -33.6191
23.1046 | -35.4312 | -33.6254
23.1107 | -35.4316 | -33.6316
23.1168 | -35.432 | -33.638
23.1229 | -35.4324 | -33.6444
23.129 | -35.4329 | -33.6507
23.1351 | -35.4333 | -33.6569
23.1413 | -35.4337 | -33.6632
23.1474 | -35.4341 | -33.6693
23.1535 | -35.4345 | -33.6756
23.1596 | -35.4349 | -33.682
23.1657 | -35.4354 | -33.6883
23.1718 | -35.4358 | -33.6946
23.1779 | -35.4362 | -33.7011
23.184 | -35.4366 | -33.7072
23.1901 | -35.437 | -33.7136
23.1962 | -35.4374 | -33.7198
23.2023 | -35.4378 | -33.7262
23.2084 | -35.4382 | -33.7325
23.2145 | -35.4387 | -33.7388
23.2206 | -35.4391 | -33.7451
23.2267 | -35.4395 | -33.7514
23.2328 | -35.4399 | -33.7576
23.2389 | -35.4403 | -33.7637
23.245 | -35.4407 | -33.7702
23.2511 | -35.4412 | -33.7762
23.2572 | -35.4416 | -33.7825
23.2633 | -35.442 | -33.7888
23.2694 | -35.4424 | -33.7951
23.2755 | -35.4428 | -33.8013
23.2817 | -35.4432 | -33.8077
23.2878 | -35.4436 | -33.8138
23.2939 | -35.444 | -33.8201
23.3 | -35.4445 | -33.8266
23.3061 | -35.4449 | -33.833
23.3122 | -35.4453 | -33.8392
23.3183 | -35.4457 | -33.8453
23.3244 | -35.4461 | -33.8513
23.3305 | -35.4465 | -33.8576
23.3366 | -35.4469 | -33.8638
23.3427 | -35.4474 | -33.8699
23.3488 | -35.4478 | -33.8762
23.3549 | -35.4482 | -33.8825
23.361 | -35.4486 | -33.8886
23.3671 | -35.449 | -33.8948
23.3732 | -35.4494 | -33.9009
23.3793 | -35.4498 | -33.907
23.3854 | -35.4502 | -33.9133
23.3915 | -35.4506 | -33.9194
23.3976 | -35.4511 | -33.9256
23.4037 | -35.4515 | -33.9318
23.4098 | -35.4519 | -33.938
23.4159 | -35.4523 | -33.9442
23.422 | -35.4527 | -33.9505
23.4282 | -35.4532 | -33.9567
23.4343 | -35.4535 | -33.963
23.4404 | -35.4539 | -33.9695
23.4465 | -35.4543 | -33.9758
23.4526 | -35.4547 | -33.9822
23.4587 | -35.4551 | -33.9884
23.4648 | -35.4555 | -33.9945
23.4709 | -35.4559 | -34.0008
23.477 | -35.4563 | -34.007
23.4831 | -35.4567 | -34.0132
23.4892 | -35.4572 | -34.0195
23.4953 | -35.4576 | -34.0256
23.5014 | -35.458 | -34.0317
23.5075 | -35.4584 | -34.0381
23.5136 | -35.4588 | -34.0444
23.5197 | -35.4592 | -34.0505
23.5258 | -35.4596 | -34.0566
23.5319 | -35.46 | -34.0626
23.538 | -35.4604 | -34.0688
23.5441 | -35.4608 | -34.0751
23.5502 | -35.4612 | -34.0811
23.5563 | -35.4616 | -34.0873
23.5624 | -35.462 | -34.0935
23.5686 | -35.4625 | -34.0997
23.5747 | -35.4628 | -34.1058
23.5808 | -35.4633 | -34.1119
23.5869 | -35.4637 | -34.1179
23.593 | -35.4641 | -34.1242
23.5991 | -35.4645 | -34.1305
23.6052 | -35.4649 | -34.1364
23.6113 | -35.4653 | -34.1427
23.6174 | -35.4657 | -34.1492
23.6235 | -35.4661 | -34.1553
23.6296 | -35.4665 | -34.1617
23.6357 | -35.4669 | -34.1679
23.6418 | -35.4672 | -34.1738
23.6479 | -35.4676 | -34.1799
23.654 | -35.4681 | -34.1861
23.6601 | -35.4685 | -34.1923
23.6662 | -35.4689 | -34.1984
23.6723 | -35.4693 | -34.2046
23.6784 | -35.4697 | -34.2108
23.6845 | -35.4701 | -34.217
23.6906 | -35.4705 | -34.223
23.6967 | -35.4709 | -34.2291
23.7028 | -35.4713 | -34.2353
23.7089 | -35.4717 | -34.2414
23.7151 | -35.4721 | -34.2476
23.7212 | -35.4725 | -34.2538
23.7273 | -35.4729 | -34.2602
23.7334 | -35.4733 | -34.2663
23.7395 | -35.4737 | -34.2725
23.7456 | -35.4741 | -34.2785
23.7517 | -35.4745 | -34.2846
23.7578 | -35.4749 | -34.2907
23.7639 | -35.4753 | -34.2969
23.77 | -35.4757 | -34.3031
23.7761 | -35.4761 | -34.3092
23.7822 | -35.4765 | -34.3151
23.7883 | -35.4769 | -34.3213
23.7944 | -35.4773 | -34.3276
23.8005 | -35.4777 | -34.3339
23.8066 | -35.4782 | -34.3401
23.8127 | -35.4786 | -34.3461
23.8188 | -35.479 | -34.3525
23.8249 | -35.4794 | -34.3587
23.831 | -35.4798 | -34.3649
23.8371 | -35.4802 | -34.371
23.8432 | -35.4806 | -34.3771
23.8493 | -35.481 | -34.3831
23.8555 | -35.4814 | -34.3892
23.8616 | -35.4818 | -34.3954
23.8677 | -35.4822 | -34.4015
23.8738 | -35.4826 | -34.4076
23.8799 | -35.483 | -34.4137
23.886 | -35.4834 | -34.4197
23.8921 | -35.4838 | -34.4258
23.8982 | -35.4842 | -34.4319
23.9043 | -35.4845 | -34.4381
23.9104 | -35.485 | -34.4443
23.9165 | -35.4854 | -34.4504
23.9226 | -35.4858 | -34.4562
23.9287 | -35.4861 | -34.4624
23.9348 | -35.4865 | -34.4687
23.9409 | -35.4869 | -34.4748
23.947 | -35.4873 | -34.4811
23.9531 | -35.4877 | -34.4868
23.9592 | -35.4881 | -34.493
23.9653 | -35.4885 | -34.499
23.9714 | -35.4889 | -34.5051
23.9775 | -35.4894 | -34.5114
23.9836 | -35.4898 | -34.5177
23.9897 | -35.4902 | -34.5237
23.9958 | -35.4906 | -34.5298
24.002 | -35.491 | -34.5359
24.0081 | -35.4914 | -34.5419
24.0142 | -35.4918 | -34.5479
24.0203 | -35.4922 | -34.5541
24.0264 | -35.4925 | -34.5601
24.0325 | -35.4929 | -34.5661
24.0386 | -35.4933 | -34.5724
24.0447 | -35.4937 | -34.5786
24.0508 | -35.4941 | -34.5845
24.0569 | -35.4945 | -34.5907
24.063 | -35.4949 | -34.5965
24.0691 | -35.4954 | -34.6027
24.0752 | -35.4958 | -34.6087
24.0813 | -35.4962 | -34.6147
24.0874 | -35.4966 | -34.6208
24.0935 | -35.497 | -34.627
24.0996 | -35.4974 | -34.6328
24.1057 | -35.4977 | -34.6389
24.1118 | -35.4981 | -34.645
24.1179 | -35.4985 | -34.651
24.124 | -35.4989 | -34.6571
24.1301 | -35.4993 | -34.6631
24.1362 | -35.4997 | -34.6692
24.1424 | -35.5001 | -34.6753
24.1485 | -35.5005 | -34.6814
24.1546 | -35.5009 | -34.6874
24.1607 | -35.5013 | -34.6935
24.1668 | -35.5017 | -34.6993
24.1729 | -35.5021 | -34.7057
24.179 | -35.5025 | -34.7119
24.1851 | -35.5029 | -34.718
24.1912 | -35.5032 | -34.7241
24.1973 | -35.5036 | -34.7302
24.2034 | -35.504 | -34.7361
24.2095 | -35.5044 | -34.7422
24.2156 | -35.5048 | -34.7483
24.2217 | -35.5052 | -34.7544
24.2278 | -35.5056 | -34.7604
24.2339 | -35.506 | -34.7663
24.24 | -35.5064 | -34.7721
24.2461 | -35.5068 | -34.7781
24.2522 | -35.5072 | -34.7842
24.2583 | -35.5076 | -34.79
24.2644 | -35.508 | -34.796
24.2705 | -35.5083 | -34.8018
24.2766 | -35.5087 | -34.8079
24.2827 | -35.5091 | -34.8139
24.2889 | -35.5095 | -34.82
24.295 | -35.5099 | -34.8259
24.3011 | -35.5103 | -34.832
24.3072 | -35.5106 | -34.8379
24.3133 | -35.511 | -34.8437
24.3194 | -35.5114 | -34.8497
24.3255 | -35.5118 | -34.8558
24.3316 | -35.5121 | -34.8618
24.3377 | -35.5125 | -34.8676
24.3438 | -35.5129 | -34.8734
24.3499 | -35.5133 | -34.8793
24.356 | -35.5137 | -34.8853
24.3621 | -35.5141 | -34.8912
24.3682 | -35.5145 | -34.8972
24.3743 | -35.5149 | -34.9032
24.3804 | -35.5153 | -34.9088
24.3865 | -35.5156 | -34.9147
24.3926 | -35.516 | -34.9208
24.3987 | -35.5164 | -34.9268
24.4048 | -35.5168 | -34.9329
24.4109 | -35.5172 | -34.9386
24.417 | -35.5175 | -34.9448
24.4231 | -35.5179 | -34.9508
24.4293 | -35.5183 | -34.9569
24.4354 | -35.5187 | -34.9629
24.4415 | -35.5191 | -34.9689
24.4476 | -35.5195 | -34.9747
24.4537 | -35.5198 | -34.9806
24.4598 | -35.5202 | -34.9864
24.4659 | -35.5206 | -34.9923
24.472 | -35.521 | -34.9983
24.4781 | -35.5214 | -35.0042
24.4842 | -35.5218 | -35.01
24.4903 | -35.5222 | -35.016
24.4964 | -35.5226 | -35.022
24.5025 | -35.5229 | -35.0279
24.5086 | -35.5233 | -35.034
24.5147 | -35.5237 | -35.0399
24.5208 | -35.5241 | -35.0453
24.5269 | -35.5245 | -35.0511
24.533 | -35.5249 | -35.0572
24.5391 | -35.5252 | -35.0631
24.5452 | -35.5256 | -35.069
24.5513 | -35.526 | -35.0748
24.5574 | -35.5264 | -35.0807
24.5635 | -35.5268 | -35.0865
24.5696 | -35.5271 | -35.0924
24.5758 | -35.5275 | -35.0983
24.5819 | -35.5279 | -35.1041
24.588 | -35.5283 | -35.1098
24.5941 | -35.5287 | -35.1156
24.6002 | -35.5291 | -35.1217
24.6063 | -35.5295 | -35.1279
24.6124 | -35.5298 | -35.1337
24.6185 | -35.5302 | -35.1397
24.6246 | -35.5306 | -35.1453
24.6307 | -35.531 | -35.1512
24.6368 | -35.5314 | -35.1571
24.6429 | -35.5318 | -35.163
24.649 | -35.5323 | -35.1688
24.6551 | -35.5326 | -35.1747
24.6612 | -35.533 | -35.1802
24.6673 | -35.5334 | -35.186
24.6734 | -35.5338 | -35.1918
24.6795 | -35.5341 | -35.1979
24.6856 | -35.5345 | -35.2039
24.6917 | -35.5349 | -35.2098
24.6978 | -35.5353 | -35.2155
24.7039 | -35.5356 | -35.2216
24.71 | -35.536 | -35.2277
24.7162 | -35.5364 | -35.2336
24.7223 | -35.5368 | -35.2394
24.7284 | -35.5372 | -35.2449
24.7345 | -35.5376 | -35.2508
24.7406 | -35.538 | -35.2566
24.7467 | -35.5383 | -35.2625
24.7528 | -35.5387 | -35.2685
24.7589 | -35.5391 | -35.2741
24.765 | -35.5395 | -35.2796
24.7711 | -35.5399 | -35.2855
24.7772 | -35.5403 | -35.2914
24.7833 | -35.5407 | -35.2973
24.7894 | -35.541 | -35.3031
24.7955 | -35.5414 | -35.3088
24.8016 | -35.5418 | -35.3144
24.8077 | -35.5422 | -35.3202
24.8138 | -35.5426 | -35.3262
24.8199 | -35.5429 | -35.3323
24.826 | -35.5433 | -35.3382
24.8321 | -35.5437 | -35.3439
24.8382 | -35.5441 | -35.3493
24.8443 | -35.5445 | -35.3551
24.8504 | -35.5448 | -35.361
24.8565 | -35.5452 | -35.3668
24.8627 | -35.5456 | -35.3727
24.8688 | -35.546 | -35.3783
24.8749 | -35.5464 | -35.3842
24.881 | -35.5468 | -35.39
24.8871 | -35.5472 | -35.3958
24.8932 | -35.5476 | -35.4018
24.8993 | -35.5479 | -35.4076
24.9054 | -35.5483 | -35.4131
24.9115 | -35.5487 | -35.419
24.9176 | -35.5491 | -35.4249
24.9237 | -35.5495 | -35.4307
24.9298 | -35.5499 | -35.4365
24.9359 | -35.5503 | -35.4425
24.942 | -35.5506 | -35.4481
24.9481 | -35.551 | -35.454
24.9542 | -35.5514 | -35.4597
24.9603 | -35.5517 | -35.4656
24.9664 | -35.5521 | -35.4715
24.9725 | -35.5525 | -35.4774
24.9786 | -35.5529 | -35.4829
24.9847 | -35.5532 | -35.4886
24.9908 | -35.5536 | -35.4945
24.9969 | -35.554 | -35.5004
25.0031 | -35.5544 | -35.5062
25.0092 | -35.5548 | -35.5119
25.0153 | -35.5552 | -35.5176
25.0214 | -35.5556 | -35.5235
25.0275 | -35.556 | -35.5293
25.0336 | -35.5564 | -35.5352
25.0397 | -35.5568 | -35.5411
25.0458 | -35.5571 | -35.5468
25.0519 | -35.5575 | -35.5525
25.058 | -35.5579 | -35.5585
25.0641 | -35.5583 | -35.5643
25.0702 | -35.5586 | -35.5701
25.0763 | -35.559 | -35.5757
25.0824 | -35.5593 | -35.5812
25.0885 | -35.5597 | -35.587
25.0946 | -35.5601 | -35.5928
25.1007 | -35.5605 | -35.5985
25.1068 | -35.5608 | -35.6041
25.1129 | -35.5612 | -35.6099
25.119 | -35.5616 | -35.6153
25.1251 | -35.5619 | -35.621
25.1312 | -35.5623 | -35.6268
25.1373 | -35.5627 | -35.6328
25.1435 | -35.5631 | -35.6386
25.1496 | -35.5635 | -35.644
25.1557 | -35.5638 | -35.6497
25.1618 | -35.5642 | -35.6554
25.1679 | -35.5646 | -35.6611
25.174 | -35.565 | -35.6668
25.1801 | -35.5653 | -35.6725
25.1862 | -35.5657 | -35.6782
25.1923 | -35.5661 | -35.6839
25.1984 | -35.5665 | -35.6897
25.2045 | -35.5668 | -35.6954
25.2106 | -35.5672 | -35.7013
25.2167 | -35.5676 | -35.7071
25.2228 | -35.5679 | -35.7125
25.2289 | -35.5683 | -35.7181
25.235 | -35.5687 | -35.7239
25.2411 | -35.569 | -35.7298
25.2472 | -35.5694 | -35.7357
25.2533 | -35.5698 | -35.7415
25.2594 | -35.5701 | -35.7468
25.2655 | -35.5705 | -35.7526
25.2716 | -35.5709 | -35.7583
25.2777 | -35.5713 | -35.7639
25.2838 | -35.5716 | -35.7697
25.29 | -35.572 | -35.7754
25.2961 | -35.5724 | -35.7807
25.3022 | -35.5728 | -35.7866
25.3083 | -35.5731 | -35.7925
25.3144 | -35.5735 | -35.7984
25.3205 | -35.5738 | -35.8041
25.3266 | -35.5742 | -35.8097
25.3327 | -35.5746 | -35.8155
25.3388 | -35.575 | -35.8213
25.3449 | -35.5753 | -35.8271
25.351 | -35.5757 | -35.8328
25.3571 | -35.5761 | -35.8385
25.3632 | -35.5765 | -35.8438
25.3693 | -35.5768 | -35.8494
25.3754 | -35.5772 | -35.8551
25.3815 | -35.5776 | -35.8608
25.3876 | -35.578 | -35.8664
25.3937 | -35.5783 | -35.8721
25.3998 | -35.5787 | -35.8774
25.4059 | -35.5791 | -35.883
25.412 | -35.5794 | -35.8886
25.4181 | -35.5798 | -35.8941
25.4242 | -35.5802 | -35.8999
25.4304 | -35.5805 | -35.9056
25.4365 | -35.5809 | -35.911
25.4426 | -35.5812 | -35.9168
25.4487 | -35.5816 | -35.9227
25.4548 | -35.582 | -35.9284
25.4609 | -35.5824 | -35.9341
25.467 | -35.5828 | -35.9395
25.4731 | -35.5831 | -35.9453
25.4792 | -35.5835 | -35.951
25.4853 | -35.5839 | -35.9568
25.4914 | -35.5843 | -35.9625
25.4975 | -35.5847 | -35.9681
25.5036 | -35.5851 | -35.9734
25.5097 | -35.5854 | -35.979
25.5158 | -35.5858 | -35.9846
25.5219 | -35.5862 | -35.9904
25.528 | -35.5865 | -35.996
25.5341 | -35.5869 | -36.0014
25.5402 | -35.5873 | -36.007
25.5463 | -35.5877 | -36.0125
25.5524 | -35.588 | -36.0183
25.5585 | -35.5884 | -36.0239
25.5646 | -35.5888 | -36.0297
25.5707 | -35.5892 | -36.0355
25.5769 | -35.5895 | -36.0408
25.583 | -35.5898 | -36.0463
25.5891 | -35.5902 | -36.0519
25.5952 | -35.5906 | -36.0577
25.6013 | -35.5909 | -36.0634
25.6074 | -35.5913 | -36.0688
25.6135 | -35.5917 | -36.0744
25.6196 | -35.5921 | -36.08
25.6257 | -35.5924 | -36.0854
25.6318 | -35.5928 | -36.091
25.6379 | -35.5932 | -36.0968
25.644 | -35.5935 | -36.1023
25.6501 | -35.5939 | -36.1079
25.6562 | -35.5942 | -36.1135
25.6623 | -35.5946 | -36.1191
25.6684 | -35.595 | -36.1247
25.6745 | -35.5954 | -36.1304
25.6806 | -35.5958 | -36.1358
25.6867 | -35.5961 | -36.1414
25.6928 | -35.5965 | -36.147
25.6989 | -35.5968 | -36.1527
25.705 | -35.5972 | -36.1583
25.7111 | -35.5976 | -36.1638
25.7173 | -35.5979 | -36.1693
25.7234 | -35.5983 | -36.1748
25.7295 | -35.5987 | -36.1805
25.7356 | -35.599 | -36.1862
25.7417 | -35.5994 | -36.1917
25.7478 | -35.5997 | -36.197
25.7539 | -35.6001 | -36.2025
25.76 | -35.6005 | -36.208
25.7661 | -35.6008 | -36.2135
25.7722 | -35.6012 | -36.219
25.7783 | -35.6015 | -36.2247
25.7844 | -35.6019 | -36.2298
25.7905 | -35.6022 | -36.2354
25.7966 | -35.6026 | -36.2412
25.8027 | -35.6029 | -36.2469
25.8088 | -35.6033 | -36.2526
25.8149 | -35.6036 | -36.2583
25.821 | -35.604 | -36.2635
25.8271 | -35.6043 | -36.2692
25.8332 | -35.6047 | -36.2749
25.8393 | -35.6051 | -36.2805
25.8454 | -35.6054 | -36.286
25.8515 | -35.6058 | -36.2915
25.8576 | -35.6061 | -36.297
25.8638 | -35.6065 | -36.3027
25.8699 | -35.6068 | -36.3082
25.876 | -35.6072 | -36.3138
25.8821 | -35.6075 | -36.3194
25.8882 | -35.6079 | -36.3248
25.8943 | -35.6082 | -36.3304
25.9004 | -35.6086 | -36.3358
25.9065 | -35.609 | -36.3413
25.9126 | -35.6093 | -36.3469
25.9187 | -35.6097 | -36.3525
25.9248 | -35.6101 | -36.3576
25.9309 | -35.6104 | -36.3631
25.937 | -35.6108 | -36.3686
25.9431 | -35.6112 | -36.374
25.9492 | -35.6115 | -36.3795
25.9553 | -35.6119 | -36.3851
25.9614 | -35.6123 | -36.3903
25.9675 | -35.6127 | -36.3958
25.9736 | -35.613 | -36.4012
25.9797 | -35.6134 | -36.4068
25.9858 | -35.6137 | -36.4122
25.9919 | -35.6141 | -36.4176
25.998 | -35.6145 | -36.423
26.0042 | -35.6148 | -36.4285
26.0103 | -35.6152 | -36.4341
26.0164 | -35.6155 | -36.4395
26.0225 | -35.6159 | -36.4449
26.0286 | -35.6162 | -36.4505
26.0347 | -35.6166 | -36.4556
26.0408 | -35.6169 | -36.4611
26.0469 | -35.6173 | -36.4664
26.053 | -35.6176 | -36.4717
26.0591 | -35.618 | -36.4772
26.0652 | -35.6183 | -36.4823
26.0713 | -35.6187 | -36.4877
26.0774 | -35.6191 | -36.4933
26.0835 | -35.6194 | -36.4988
26.0896 | -35.6198 | -36.5045
26.0957 | -35.6201 | -36.5103
26.1018 | -35.6205 | -36.5155
26.1079 | -35.6208 | -36.521
26.114 | -35.6212 | -36.5265
26.1201 | -35.6215 | -36.5319
26.1262 | -35.6219 | -36.5374
26.1323 | -35.6223 | -36.543
26.1384 | -35.6226 | -36.5485
26.1445 | -35.623 | -36.554
26.1507 | -35.6233 | -36.5597
26.1568 | -35.6237 | -36.5651
26.1629 | -35.6241 | -36.5705
26.169 | -35.6244 | -36.5759
26.1751 | -35.6247 | -36.5812
26.1812 | -35.6251 | -36.5865
26.1873 | -35.6254 | -36.5921
26.1934 | -35.6258 | -36.5973
26.1995 | -35.6262 | -36.6028
26.2056 | -35.6265 | -36.6081
26.2117 | -35.6268 | -36.6136
26.2178 | -35.6272 | -36.619
26.2239 | -35.6275 | -36.6245
26.23 | -35.6279 | -36.6301
26.2361 | -35.6282 | -36.6358
26.2422 | -35.6286 | -36.6412
26.2483 | -35.6289 | -36.6467
26.2544 | -35.6292 | -36.6521
26.2605 | -35.6296 | -36.6577
26.2666 | -35.63 | -36.6634
26.2727 | -35.6303 | -36.6686
26.2788 | -35.6307 | -36.6738
26.2849 | -35.6311 | -36.6792
26.2911 | -35.6314 | -36.6847
26.2972 | -35.6318 | -36.6903
26.3033 | -35.6321 | -36.6957
26.3094 | -35.6325 | -36.7011
26.3155 | -35.6328 | -36.7062
26.3216 | -35.6331 | -36.7116
26.3277 | -35.6335 | -36.717
26.3338 | -35.6338 | -36.7224
26.3399 | -35.6342 | -36.7276
26.346 | -35.6345 | -36.7326
26.3521 | -35.6349 | -36.738
26.3582 | -35.6352 | -36.7433
26.3643 | -35.6355 | -36.7489
26.3704 | -35.6359 | -36.7542
26.3765 | -35.6362 | -36.7597
26.3826 | -35.6365 | -36.765
26.3887 | -35.6369 | -36.7703
26.3948 | -35.6373 | -36.7757
26.4009 | -35.6376 | -36.7811
26.407 | -35.638 | -36.7867
26.4131 | -35.6384 | -36.7921
26.4192 | -35.6387 | -36.7972
26.4253 | -35.6391 | -36.8027
26.4314 | -35.6394 | -36.808
26.4376 | -35.6398 | -36.8135
26.4437 | -35.6401 | -36.8188
26.4498 | -35.6404 | -36.8247
26.4559 | -35.6408 | -36.8297
26.462 | -35.6411 | -36.8352
26.4681 | -35.6415 | -36.8406
26.4742 | -35.6418 | -36.846
26.4803 | -35.6422 | -36.8515
26.4864 | -35.6425 | -36.8568
26.4925 | -35.6429 | -36.8622
26.4986 | -35.6433 | -36.8678
26.5047 | -35.6436 | -36.8731
26.5108 | -35.644 | -36.8783
26.5169 | -35.6443 | -36.8838
26.523 | -35.6447 | -36.8888
26.5291 | -35.645 | -36.8943
26.5352 | -35.6453 | -36.8998
26.5413 | -35.6457 | -36.9049
26.5474 | -35.6461 | -36.9103
26.5535 | -35.6464 | -36.9158
26.5596 | -35.6468 | -36.921
26.5657 | -35.6471 | -36.9265
26.5718 | -35.6475 | -36.932
26.578 | -35.6478 | -36.9375
26.5841 | -35.6482 | -36.9429
26.5902 | -35.6485 | -36.9483
26.5963 | -35.6489 | -36.9536
26.6024 | -35.6492 | -36.9587
26.6085 | -35.6496 | -36.9639
26.6146 | -35.6499 | -36.9693
26.6207 | -35.6503 | -36.9747
26.6268 | -35.6507 | -36.9801
26.6329 | -35.651 | -36.9851
26.639 | -35.6513 | -36.9907
26.6451 | -35.6516 | -36.996
26.6512 | -35.652 | -37.0015
26.6573 | -35.6524 | -37.0068
26.6634 | -35.6527 | -37.0121
26.6695 | -35.6531 | -37.0176
26.6756 | -35.6534 | -37.023
26.6817 | -35.6537 | -37.0284
26.6878 | -35.6541 | -37.0337
26.6939 | -35.6544 | -37.039
26.7 | -35.6547 | -37.0443
26.7061 | -35.6551 | -37.0496
26.7122 | -35.6554 | -37.055
26.7183 | -35.6557 | -37.0604
26.7245 | -35.6561 | -37.0656
26.7306 | -35.6564 | -37.071
26.7367 | -35.6567 | -37.0761
26.7428 | -35.6571 | -37.0813
26.7489 | -35.6574 | -37.0868
26.755 | -35.6578 | -37.0921
26.7611 | -35.6581 | -37.0974
26.7672 | -35.6584 | -37.103
26.7733 | -35.6588 | -37.1081
26.7794 | -35.6591 | -37.1134
26.7855 | -35.6595 | -37.1187
26.7916 | -35.6598 | -37.1241
26.7977 | -35.6602 | -37.1293
26.8038 | -35.6605 | -37.1348
26.8099 | -35.6609 | -37.14
26.816 | -35.6612 | -37.1454
26.8221 | -35.6615 | -37.1508
26.8282 | -35.6619 | -37.1562
26.8343 | -35.6622 | -37.1616
26.8404 | -35.6625 | -37.1667
26.8465 | -35.6629 | -37.172
26.8526 | -35.6632 | -37.1772
26.8587 | -35.6635 | -37.1824
26.8649 | -35.6639 | -37.1879
26.871 | -35.6642 | -37.1931
26.8771 | -35.6646 | -37.1984
26.8832 | -35.6649 | -37.2036
26.8893 | -35.6653 | -37.2087
26.8954 | -35.6656 | -37.2142
26.9015 | -35.666 | -37.2194
26.9076 | -35.6663 | -37.2248
26.9137 | -35.6666 | -37.23
26.9198 | -35.667 | -37.2354
26.9259 | -35.6673 | -37.2408
26.932 | -35.6677 | -37.246
26.9381 | -35.668 | -37.2514
26.9442 | -35.6684 | -37.2566
26.9503 | -35.6687 | -37.2621
26.9564 | -35.6691 | -37.2674
26.9625 | -35.6695 | -37.2728
26.9686 | -35.6698 | -37.2781
26.9747 | -35.6701 | -37.2834
26.9808 | -35.6705 | -37.2884
26.9869 | -35.6708 | -37.2936
26.993 | -35.6711 | -37.299
26.9991 | -35.6715 | -37.3042
27.0052 | -35.6718 | -37.3096
27.0114 | -35.6721 | -37.3148
27.0175 | -35.6724 | -37.3201
27.0236 | -35.6728 | -37.3253
27.0297 | -35.6732 | -37.3305
27.0358 | -35.6735 | -37.3358
27.0419 | -35.6739 | -37.3413
27.048 | -35.6742 | -37.3466
27.0541 | -35.6745 | -37.3516
27.0602 | -35.6749 | -37.3569
27.0663 | -35.6752 | -37.3623
27.0724 | -35.6756 | -37.3674
27.0785 | -35.6759 | -37.3728
27.0846 | -35.6763 | -37.3778
27.0907 | -35.6766 | -37.3829
27.0968 | -35.677 | -37.3884
27.1029 | -35.6773 | -37.3936
27.109 | -35.6777 | -37.3986
27.1151 | -35.678 | -37.404
27.1212 | -35.6783 | -37.4092
27.1273 | -35.6787 | -37.4146
27.1334 | -35.679 | -37.4202
27.1395 | -35.6794 | -37.4256
27.1456 | -35.6797 | -37.4307
27.1518 | -35.6801 | -37.436
27.1579 | -35.6804 | -37.4412
27.164 | -35.6807 | -37.4463
27.1701 | -35.6811 | -37.4517
27.1762 | -35.6814 | -37.4572
27.1823 | -35.6817 | -37.4625
27.1884 | -35.6821 | -37.4677
27.1945 | -35.6824 | -37.473
27.2006 | -35.6827 | -37.4782
27.2067 | -35.6831 | -37.4833
27.2128 | -35.6834 | -37.4886
27.2189 | -35.6838 | -37.4939
27.225 | -35.6841 | -37.4992
27.2311 | -35.6845 | -37.5043
27.2372 | -35.6848 | -37.5095
27.2433 | -35.6851 | -37.5147
27.2494 | -35.6855 | -37.5199
27.2555 | -35.6858 | -37.5252
27.2616 | -35.6861 | -37.5302
27.2677 | -35.6865 | -37.5356
27.2738 | -35.6868 | -37.541
27.2799 | -35.6871 | -37.5464
27.286 | -35.6874 | -37.5514
27.2921 | -35.6878 | -37.5566
27.2983 | -35.6881 | -37.5618
27.3044 | -35.6884 | -37.5672
27.3105 | -35.6887 | -37.5723
27.3166 | -35.6891 | -37.5774
27.3227 | -35.6894 | -37.5825
27.3288 | -35.6898 | -37.5876
27.3349 | -35.6901 | -37.5927
27.341 | -35.6904 | -37.5977
27.3471 | -35.6908 | -37.6028
27.3532 | -35.6911 | -37.6079
27.3593 | -35.6915 | -37.6131
27.3654 | -35.6918 | -37.6184
27.3715 | -35.6921 | -37.6235
27.3776 | -35.6924 | -37.6287
27.3837 | -35.6928 | -37.634
27.3898 | -35.6931 | -37.6391
27.3959 | -35.6934 | -37.6445
27.402 | -35.6938 | -37.6494
27.4081 | -35.6941 | -37.6547
27.4142 | -35.6945 | -37.6599
27.4203 | -35.6948 | -37.6653
27.4264 | -35.6951 | -37.6705
27.4325 | -35.6954 | -37.6758
27.4387 | -35.6958 | -37.6811
27.4448 | -35.6961 | -37.6863
27.4509 | -35.6964 | -37.6916
27.457 | -35.6968 | -37.6969
27.4631 | -35.6971 | -37.7021
27.4692 | -35.6974 | -37.7073
27.4753 | -35.6977 | -37.7122
27.4814 | -35.698 | -37.7175
27.4875 | -35.6983 | -37.7228
27.4936 | -35.6987 | -37.728
27.4997 | -35.699 | -37.7332
27.5058 | -35.6993 | -37.7385
27.5119 | -35.6997 | -37.7436
27.518 | -35.7 | -37.7487
27.5241 | -35.7004 | -37.754
27.5302 | -35.7007 | -37.7592
27.5363 | -35.701 | -37.7645
27.5424 | -35.7013 | -37.7694
27.5485 | -35.7016 | -37.7745
27.5546 | -35.7019 | -37.7796
27.5607 | -35.7023 | -37.7849
27.5668 | -35.7026 | -37.79
27.5729 | -35.7029 | -37.7951
27.5791 | -35.7032 | -37.8001
27.5852 | -35.7036 | -37.8054
27.5913 | -35.7039 | -37.8106
27.5974 | -35.7042 | -37.8157
27.6035 | -35.7046 | -37.821
27.6096 | -35.7049 | -37.8261
27.6157 | -35.7052 | -37.8313
27.6218 | -35.7056 | -37.8366
27.6279 | -35.7059 | -37.8418
27.634 | -35.7062 | -37.8472
27.6401 | -35.7065 | -37.8523
27.6462 | -35.7068 | -37.8573
27.6523 | -35.7072 | -37.8624
27.6584 | -35.7075 | -37.8674
27.6645 | -35.7078 | -37.8726
27.6706 | -35.7081 | -37.8777
27.6767 | -35.7084 | -37.8827
27.6828 | -35.7087 | -37.8878
27.6889 | -35.7091 | -37.8929
27.695 | -35.7094 | -37.8982
27.7011 | -35.7097 | -37.9035
27.7072 | -35.71 | -37.9087
27.7133 | -35.7104 | -37.9139
27.7194 | -35.7107 | -37.919
27.7256 | -35.7111 | -37.9243
27.7317 | -35.7114 | -37.9296
27.7378 | -35.7117 | -37.9347
27.7439 | -35.712 | -37.9402
27.75 | -35.7124 | -37.9454
27.7561 | -35.7127 | -37.9507
27.7622 | -35.713 | -37.9559
27.7683 | -35.7134 | -37.9611
27.7744 | -35.7137 | -37.966
27.7805 | -35.7141 | -37.9711
27.7866 | -35.7144 | -37.9765
27.7927 | -35.7147 | -37.9813
27.7988 | -35.7151 | -37.9864
27.8049 | -35.7154 | -37.9913
27.811 | -35.7157 | -37.9964
27.8171 | -35.7161 | -38.0018
27.8232 | -35.7164 | -38.007
27.8293 | -35.7168 | -38.0122
27.8354 | -35.7171 | -38.0174
27.8415 | -35.7174 | -38.0225
27.8476 | -35.7177 | -38.0277
27.8537 | -35.7181 | -38.0329
27.8598 | -35.7184 | -38.038
27.866 | -35.7187 | -38.0432
27.8721 | -35.719 | -38.0484
27.8782 | -35.7193 | -38.0538
27.8843 | -35.7196 | -38.0591
27.8904 | -35.7199 | -38.0641
27.8965 | -35.7202 | -38.069
27.9026 | -35.7206 | -38.0741
27.9087 | -35.7209 | -38.0793
27.9148 | -35.7212 | -38.0846
27.9209 | -35.7215 | -38.0898
27.927 | -35.7218 | -38.0949
27.9331 | -35.7221 | -38.0999
27.9392 | -35.7224 | -38.105
27.9453 | -35.7228 | -38.1101
27.9514 | -35.7231 | -38.1154
27.9575 | -35.7234 | -38.1206
27.9636 | -35.7237 | -38.126
27.9697 | -35.7241 | -38.1308
27.9758 | -35.7244 | -38.1362
27.9819 | -35.7247 | -38.1416
27.988 | -35.7251 | -38.1468
27.9941 | -35.7254 | -38.152
28.0002 | -35.7257 | -38.157
28.0063 | -35.726 | -38.1621
28.0125 | -35.7263 | -38.1671
28.0186 | -35.7266 | -38.1722
28.0247 | -35.7269 | -38.1774
28.0308 | -35.7272 | -38.1823
28.0369 | -35.7275 | -38.1872
28.043 | -35.7278 | -38.1924
28.0491 | -35.7282 | -38.1977
28.0552 | -35.7285 | -38.2029
28.0613 | -35.7288 | -38.2079
28.0674 | -35.7291 | -38.2129
28.0735 | -35.7294 | -38.2178

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4525 | -43.0529
21.4908 | -35.5106 | -35.0038
16.4948 | -34.9195 | -27.5955
12.6602 | -34.5588 | -21.3173
9.7171 | -34.3386 | -16.413
7.4582 | -34.1972 | -12.4373
5.7244 | -34.0932 | -9.2144
4.3936 | -34.0234 | -6.6239
3.3722 | -34.02 | -4.5797
2.5883 | -34.0061 | -3.0294
1.9866 | -34.0033 | -1.7952
1.5248 | -33.981 | -0.80176
1.1703 | -33.9921 | -0.15406
1 | -33.9911 | 0.1654
0.89822 | -33.9802 | 0.43592
0.68941 | -33.9652 | 1.0143
0.52919 | -33.9598 | 1.4664
0.40612 | -33.9584 | 2.0727
0.31171 | -33.9736 | 2.5261
0.23928 | -33.9643 | 3.2514
0.18358 | -33.9656 | 3.6804
0.14091 | -33.9768 | 4.5578
0.10819 | -34 | 5.7141
0.083017 | -34.0348 | 7.1915
0.063705 | -34.0738 | 9.0648
0.048923 | -34.1417 | 11.5575
0.037479 | -34.2398 | 14.7454
0.028801 | -34.4194 | 18.7684
0.022078 | -34.723 | 24.0656
0.017023 | -35.1921 | 30.4623
0.013018 | -35.9956 | 38.1458
0.010014 | -37.1156 | 46.018

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003726 | -3.4785
0.007452 | -28.0103
0.011178 | -35.7426
0.014904 | -40.2603
0.01863 | -42.2857
0.022356 | -44.9802
0.026082 | -47.3941
0.029808 | -49.1258
0.033534 | -50.7768
0.03726 | -52.0595
0.040986 | -53.4794
0.044711 | -54.4458
0.048437 | -55.8544
0.052163 | -56.1891
0.055889 | -56.4043
0.059615 | -57.1667
0.063341 | -57.1133
0.067067 | -56.7087
0.070793 | -57.2593
0.074519 | -58.7891
0.078245 | -59.2677
0.081971 | -60.8902
0.085697 | -61.1223
0.089423 | -61.5233
0.093149 | -62.2248
0.096875 | -62.7558
0.1006 | -63.5886
0.10433 | -63.1935
0.10805 | -63.2856
0.11178 | -63.9259
0.1155 | -64.694
0.11923 | -64.7846
0.12296 | -64.8388
0.12668 | -65.1002
0.13041 | -65.6085
0.13413 | -65.9628
0.13786 | -66.1161
0.14159 | -66.2871
0.14531 | -66.5969
0.14904 | -66.9736
0.15276 | -66.3459
0.15649 | -65.5444
0.16022 | -65.9254
0.16394 | -66.7643
0.16767 | -67.3804
0.17139 | -67.9227
0.17512 | -68.3174
0.17885 | -68.0979
0.18257 | -68.6809
0.1863 | -68.6735
0.19002 | -68.3589
0.19375 | -67.6398
0.19748 | -67.5215
0.2012 | -68.1555
0.20493 | -69.123
0.20865 | -69.0641
0.21238 | -69.6138
0.2161 | -70.2133
0.21983 | -70.7542
0.22356 | -70.4456
0.22728 | -69.8378
0.23101 | -70.0391
0.23473 | -69.6867
0.23846 | -69.4844
0.24219 | -70.5384
0.24591 | -70.8221
0.24964 | -70.9791
0.25337 | -71.1491
0.25709 | -70.5434
0.26082 | -70.8462
0.26454 | -72.5679
0.26827 | -72.4301
0.27199 | -72.653
0.27572 | -72.6429
0.27945 | -72.4475
0.28317 | -71.7375
0.2869 | -71.8512
0.29063 | -72.6917
0.29435 | -72.3871
0.29808 | -72.2688
0.3018 | -73.1552
0.30553 | -73.1086
0.30925 | -72.7515
0.31298 | -73.0773
0.31671 | -73.05
0.32043 | -72.7806
0.32416 | -73.6699
0.32788 | -74.6275
0.33161 | -73.7878
0.33534 | -72.9673
0.33906 | -73.7597
0.34279 | -74.7405
0.34651 | -74.3132
0.35024 | -74.636
0.35397 | -74.4061
0.35769 | -74.2512
0.36142 | -75.5013
0.36514 | -75.4737
0.36887 | -74.7883
0.3726 | -74.3686
0.37632 | -74.8608
0.38005 | -75.7349
0.38377 | -75.44
0.3875 | -74.8784
0.39122 | -74.9362
0.39495 | -75.4617
0.39868 | -76.9336
0.4024 | -76.3745
0.40613 | -75.3015
0.40986 | -75.1348
0.41358 | -75.4007
0.41731 | -75.8458
0.42103 | -75.8124
0.42476 | -75.2292
0.42849 | -75.3941
0.43221 | -75.7123
0.43594 | -75.5757
0.43966 | -75.2903
0.44339 | -76.0075
0.44711 | -76.2285
0.45084 | -75.9777
0.45457 | -76.4487
0.45829 | -75.9039
0.46202 | -76.2441
0.46574 | -76.7171
0.46947 | -76.6421
0.4732 | -76.8538
0.47692 | -76.6726
0.48065 | -76.7507
0.48437 | -77.439
0.4881 | -76.5285
0.49183 | -76.1352
0.49555 | -76.6313
0.49928 | -76.7376
0.503 | -76.7645
0.50673 | -77.3759
0.51046 | -78.4556
0.51418 | -78.992
0.51791 | -78.9411
0.52163 | -78.1854
0.52536 | -77.1543
0.52909 | -77.618
0.53281 | -77.7069
0.53654 | -77.5977
0.54026 | -77.476
0.54399 | -77.7682
0.54772 | -78.0614
0.55144 | -77.8429
0.55517 | -78.0491
0.55889 | -78.6884
0.56262 | -78.7444
0.56634 | -78.4474
0.57007 | -78.7556
0.5738 | -79.344
0.57752 | -78.9481
0.58125 | -78.1669
0.58498 | -78.1348
0.5887 | -78.7035
0.59243 | -79.2241
0.59615 | -79.3117
0.59988 | -79.155
0.6036 | -78.887
0.60733 | -78.8388
0.61106 | -79.3235
0.61478 | -80.1907
0.61851 | -80.5416
0.62223 | -80.1266
0.62596 | -79.7149
0.62969 | -79.6718
0.63341 | -79.6383
0.63714 | -79.6111
0.64086 | -79.6998
0.64459 | -80.0325
0.64832 | -80.144
0.65204 | -79.7317
0.65577 | -79.4913
0.65949 | -79.5923
0.66322 | -79.5313
0.66695 | -79.2017
0.67067 | -79.1628
0.6744 | -79.5641
0.67812 | -79.9836
0.68185 | -80.3063
0.68558 | -80.8876
0.6893 | -81.655
0.69303 | -82.0875
0.69675 | -81.7849
0.70048 | -81.0962
0.70421 | -80.9505
0.70793 | -81.2335
0.71166 | -81.4125
0.71538 | -81.505
0.71911 | -81.3894
0.72284 | -81.0014
0.72656 | -80.6388
0.73029 | -80.567
0.73401 | -80.5731
0.73774 | -80.5643
0.74147 | -80.7138
0.74519 | -81.0436
0.74892 | -80.9893
0.75264 | -80.6491
0.75637 | -80.6217
0.76009 | -81.0913
0.76382 | -81.7415
0.76755 | -81.9211
0.77127 | -81.7798
0.775 | -81.8912
0.77873 | -82.5202
0.78245 | -83.1191
0.78618 | -83.0255
0.7899 | -82.0366
0.79363 | -81.4075
0.79736 | -81.9102
0.80108 | -82.5245
0.80481 | -82.3716
0.80853 | -81.9532
0.81226 | -81.7218
0.81598 | -81.608
0.81971 | -81.4404
0.82344 | -81.2658
0.82716 | -81.0738
0.83089 | -81.1098
0.83461 | -81.3007
0.83834 | -81.6077
0.84207 | -82.2034
0.84579 | -82.6128
0.84952 | -82.7902
0.85324 | -83.2079
0.85697 | -83.4417
0.8607 | -82.9423
0.86442 | -82.0477
0.86815 | -81.6458
0.87187 | -81.9241
0.8756 | -82.6116
0.87933 | -83.0427
0.88305 | -83.0506
0.88678 | -82.929
0.8905 | -82.9517
0.89423 | -83.3018
0.89795 | -83.6707
0.90168 | -83.7492
0.90541 | -83.6109
0.90913 | -83.3941
0.91286 | -83.4257
0.91658 | -83.7655
0.92031 | -84.2265
0.92404 | -84.3056
0.92776 | -83.9135
0.93149 | -83.7003
0.93522 | -83.4803
0.93894 | -83.1938
0.94267 | -83.1277
0.94639 | -83.2044
0.95012 | -83.2815
0.95385 | -83.465
0.95757 | -83.7709
0.9613 | -84.2075
0.96502 | -84.3076
0.96875 | -83.8911
0.97247 | -83.5181
0.9762 | -83.7263
0.97993 | -84.1421
0.98365 | -84.1186
0.98738 | -83.9129
0.9911 | -83.8635
0.99483 | -84.0107
0.99856 | -84.2949
1.0023 | -84.3209
1.006 | -83.9755
1.0097 | -83.8683
1.0135 | -84.0505
1.0172 | -84.2473
1.0209 | -84.2622
1.0246 | -84.4555
1.0284 | -84.7678
1.0321 | -84.7834
1.0358 | -84.4908
1.0395 | -83.9939
1.0433 | -83.5025
1.047 | -83.3289
1.0507 | -83.414
1.0544 | -83.4223
1.0582 | -83.2157
1.0619 | -83.129
1.0656 | -83.3093
1.0693 | -83.5193
1.0731 | -83.8171
1.0768 | -84.0906
1.0805 | -84.2193
1.0843 | -84.1633
1.088 | -84.0351
1.0917 | -83.9665
1.0954 | -84.0247
1.0992 | -84.2043
1.1029 | -84.5029
1.1066 | -84.9039
1.1103 | -85.2914
1.1141 | -85.5602
1.1178 | -85.721
1.1215 | -85.7577
1.1252 | -85.6473
1.129 | -85.41
1.1327 | -85.0885
1.1364 | -84.7683
1.1401 | -84.5286
1.1439 | -84.4511
1.1476 | -84.5458
1.1513 | -84.7161
1.155 | -84.8996
1.1588 | -85.1179
1.1625 | -85.4245
1.1662 | -85.7513
1.17 | -85.8766
1.1737 | -85.7414
1.1774 | -85.4762
1.1811 | -85.235
1.1849 | -85.1406
1.1886 | -85.1095
1.1923 | -85.0213
1.196 | -84.9441
1.1998 | -85.016
1.2035 | -85.2997
1.2072 | -85.6573
1.2109 | -85.8441
1.2147 | -85.8133
1.2184 | -85.668
1.2221 | -85.5709
1.2258 | -85.6703
1.2296 | -85.9363
1.2333 | -86.2107
1.237 | -86.3381
1.2407 | -86.2186
1.2445 | -85.9814
1.2482 | -85.8899
1.2519 | -86.0306
1.2556 | -86.2568
1.2594 | -86.2978
1.2631 | -86.0904
1.2668 | -85.9197
1.2706 | -86.0071
1.2743 | -86.2856
1.278 | -86.5325
1.2817 | -86.5725
1.2855 | -86.4108
1.2892 | -86.1638
1.2929 | -85.8354
1.2966 | -85.502
1.3004 | -85.3205
1.3041 | -85.2528
1.3078 | -85.2312
1.3115 | -85.3299
1.3153 | -85.5182
1.319 | -85.6358
1.3227 | -85.7248
1.3264 | -85.9859
1.3302 | -86.3837
1.3339 | -86.6439
1.3376 | -86.5772
1.3413 | -86.2352
1.3451 | -85.8242
1.3488 | -85.5633
1.3525 | -85.5591
1.3562 | -85.7835
1.36 | -86.1571
1.3637 | -86.5362
1.3674 | -86.7501
1.3712 | -86.7907
1.3749 | -86.7194
1.3786 | -86.4945
1.3823 | -86.1166
1.3861 | -85.7666
1.3898 | -85.6962
1.3935 | -86.0246
1.3972 | -86.5382
1.401 | -86.8351
1.4047 | -86.8785
1.4084 | -86.9461
1.4121 | -87.08
1.4159 | -87.0835
1.4196 | -86.9843
1.4233 | -87.0015
1.427 | -87.0834
1.4308 | -87.0152
1.4345 | -86.7691
1.4382 | -86.5208
1.4419 | -86.4716
1.4457 | -86.6644
1.4494 | -86.9458
1.4531 | -87.1234
1.4568 | -87.1963
1.4606 | -87.2786
1.4643 | -87.3813
1.468 | -87.4274
1.4718 | -87.3667
1.4755 | -87.2818
1.4792 | -87.2428
1.4829 | -87.1519
1.4867 | -87.0225
1.4904 | -87.0378
1.4941 | -87.2069
1.4978 | -87.3187
1.5016 | -87.1567
1.5053 | -86.8397
1.509 | -86.721
1.5127 | -86.9114
1.5165 | -87.2632
1.5202 | -87.6336
1.5239 | -87.8849
1.5276 | -87.8929
1.5314 | -87.7381
1.5351 | -87.5872
1.5388 | -87.5384
1.5425 | -87.6197
1.5463 | -87.6575
1.55 | -87.5184
1.5537 | -87.3706
1.5575 | -87.4079
1.5612 | -87.7576
1.5649 | -88.3761
1.5686 | -88.8795
1.5724 | -88.9129
1.5761 | -88.6217
1.5798 | -88.3559
1.5835 | -88.2189
1.5873 | -88.0481
1.591 | -87.7878
1.5947 | -87.6614
1.5984 | -87.8486
1.6022 | -88.2415
1.6059 | -88.5288
1.6096 | -88.2313
1.6133 | -87.9039
1.6171 | -87.4953
1.6208 | -87.1792
1.6245 | -87.0454
1.6282 | -87.1043
1.632 | -87.3108
1.6357 | -87.5793
1.6394 | -87.8019
1.6431 | -87.91
1.6469 | -87.911
1.6506 | -87.8417
1.6543 | -87.7177
1.6581 | -87.574
1.6618 | -87.4631
1.6655 | -87.4218
1.6692 | -87.4878
1.673 | -87.6981
1.6767 | -88.048
1.6804 | -88.4489
1.6841 | -88.7449
1.6879 | -88.8422
1.6916 | -88.7873
1.6953 | -88.6734
1.699 | -88.5322
1.7028 | -88.3426
1.7065 | -88.1769
1.7102 | -88.1064
1.7139 | -88.1136
1.7177 | -88.1296
1.7214 | -88.1322
1.7251 | -88.1711
1.7288 | -88.2461
1.7326 | -88.2725
1.7363 | -88.1954
1.74 | -88.0521
1.7437 | -87.9588
1.7475 | -88.0156
1.7512 | -88.2709
1.7549 | -88.6702
1.7587 | -89.0279
1.7624 | -89.1625
1.7661 | -89.0533
1.7698 | -88.8085
1.7736 | -88.5496
1.7773 | -88.3596
1.781 | -88.2717
1.7847 | -88.2845
1.7885 | -88.4143
1.7922 | -88.6831
1.7959 | -89.0316
1.7996 | -89.3072
1.8034 | -89.3758
1.8071 | -89.246
1.8108 | -89.0389
1.8145 | -88.891
1.8183 | -88.8737
1.822 | -88.9471
1.8257 | -88.9896
1.8294 | -88.8982
1.8332 | -88.7075
1.8369 | -88.569
1.8406 | -88.5841
1.8443 | -88.7248
1.8481 | -88.8556
1.8518 | -88.8615
1.8555 | -88.7698
1.8593 | -88.7221
1.863 | -88.823
1.8667 | -89.0601
1.8704 | -89.3537
1.8742 | -89.62
1.8779 | -89.7644
1.8816 | -89.7058
1.8853 | -89.4603
1.8891 | -89.1587
1.8928 | -88.911
1.8965 | -88.7367
1.9002 | -88.6309
1.904 | -88.6021
1.9077 | -88.6654
1.9114 | -88.8125
1.9151 | -89.0175
1.9189 | -89.2447
1.9226 | -89.4639
1.9263 | -89.6504
1.93 | -89.789
1.9338 | -89.8528
1.9375 | -89.8421
1.9412 | -89.7842
1.9449 | -89.726
1.9487 | -89.7194
1.9524 | -89.7591
1.9561 | -89.7653
1.9599 | -89.6732
1.9636 | -89.5051
1.9673 | -89.3834
1.971 | -89.4496
1.9748 | -89.7205
1.9785 | -90.0185
1.9822 | -90.0516
1.9859 | -89.7133
1.9897 | -89.2119
1.9934 | -88.8207
1.9971 | -88.6478
2.0008 | -88.6598
2.0046 | -88.7794
2.0083 | -88.9724
2.012 | -89.1918
2.0157 | -89.308
2.0195 | -89.2326
2.0232 | -89.0263
2.0269 | -88.863
2.0306 | -88.9004
2.0344 | -89.1684
2.0381 | -89.5793
2.0418 | -90.0101
2.0455 | -90.4289
2.0493 | -90.8753
2.053 | -91.2831
2.0567 | -91.4971
2.0605 | -91.3888
2.0642 | -91.0109
2.0679 | -90.5401
2.0716 | -90.1002
2.0754 | -89.7501
2.0791 | -89.5442
2.0828 | -89.521
2.0865 | -89.6696
2.0903 | -89.9341
2.094 | -90.2631
2.0977 | -90.5853
2.1014 | -90.7749
2.1052 | -90.7682
2.1089 | -90.6376
2.1126 | -90.5423
2.1163 | -90.5957
2.1201 | -90.8028
2.1238 | -91.0743
2.1275 | -91.2697
2.1312 | -91.2961
2.135 | -91.154
2.1387 | -90.9122
2.1424 | -90.7092
2.1462 | -90.6368
2.1499 | -90.7128
2.1536 | -90.7389
2.1573 | -90.7558
2.1611 | -90.7643
2.1648 | -90.7893
2.1685 | -90.8362
2.1722 | -90.8937
2.176 | -90.9454
2.1797 | -90.9916
2.1834 | -91.0464
2.1871 | -91.1086
2.1909 | -91.1839
2.1946 | -91.2496
2.1983 | -91.2736
2.202 | -91.2332
2.2058 | -91.1516
2.2095 | -91.0797
2.2132 | -91.0533
2.2169 | -91.0577
2.2207 | -91.0321
2.2244 | -90.9322
2.2281 | -90.7603
2.2318 | -90.5597
2.2356 | -90.4056
2.2393 | -90.3536
2.243 | -90.4267
2.2468 | -90.5971
2.2505 | -90.7998
2.2542 | -90.9704
2.2579 | -91.0728
2.2617 | -91.1101
2.2654 | -91.1044
2.2691 | -91.0607
2.2728 | -90.978
2.2766 | -90.8595
2.2803 | -90.7309
2.284 | -90.6223
2.2877 | -90.5398
2.2915 | -90.4646
2.2952 | -90.4027
2.2989 | -90.3806
2.3026 | -90.4169
2.3064 | -90.5223
2.3101 | -90.6992
2.3138 | -90.9215
2.3175 | -91.1464
2.3213 | -91.3298
2.325 | -91.4301
2.3287 | -91.4659
2.3324 | -91.4782
2.3362 | -91.4656
2.3399 | -91.4404
2.3436 | -91.4394
2.3474 | -91.5
2.3511 | -91.6243
2.3548 | -91.7748
2.3585 | -91.8776
2.3623 | -91.8768
2.366 | -91.8031
2.3697 | -91.7028
2.3734 | -91.6147
2.3772 | -91.5312
2.3809 | -91.4162
2.3846 | -91.2567
2.3883 | -91.087
2.3921 | -90.9627
2.3958 | -90.9041
2.3995 | -90.905
2.4032 | -90.9495
2.407 | -91.0272
2.4107 | -91.1332
2.4144 | -91.2386
2.4181 | -91.3201
2.4219 | -91.3431
2.4256 | -91.3044
2.4293 | -91.2209
2.433 | -91.1175
2.4368 | -91.0306
2.4405 | -91.0086
2.4442 | -91.0743
2.448 | -91.2209
2.4517 | -91.3898
2.4554 | -91.5126
2.4591 | -91.5574
2.4629 | -91.5635
2.4666 | -91.5643
2.4703 | -91.5529
2.474 | -91.51
2.4778 | -91.418
2.4815 | -91.2803
2.4852 | -91.139
2.4889 | -91.0463
2.4927 | -91.0593
2.4964 | -91.1821
2.5001 | -91.3601
2.5038 | -91.5166
2.5076 | -91.5846
2.5113 | -91.539
2.515 | -91.4012
2.5187 | -91.2441
2.5225 | -91.1532
2.5262 | -91.1913
2.5299 | -91.3662
2.5337 | -91.6352
2.5374 | -91.8956
2.5411 | -92.0469
2.5448 | -92.073
2.5486 | -92.0404
2.5523 | -92.0269
2.556 | -92.0647
2.5597 | -92.1399
2.5635 | -92.2149
2.5672 | -92.2662
2.5709 | -92.2835
2.5746 | -92.2509
2.5784 | -92.1653
2.5821 | -92.0113
2.5858 | -91.7751
2.5895 | -91.4923
2.5933 | -91.2347
2.597 | -91.0646
2.6007 | -91.0086
2.6044 | -91.0662
2.6082 | -91.1887
2.6119 | -91.3051
2.6156 | -91.3474
2.6193 | -91.3046
2.6231 | -91.2283
2.6268 | -91.1719
2.6305 | -91.1465
2.6343 | -91.1283
2.638 | -91.1099
2.6417 | -91.1083
2.6454 | -91.1297
2.6492 | -91.1388
2.6529 | -91.0751
2.6566 | -90.9606
2.6603 | -90.8752
2.6641 | -90.8839
2.6678 | -91.0068
2.6715 | -91.2075
2.6752 | -91.4525
2.679 | -91.7201
2.6827 | -91.9631
2.6864 | -92.0376
2.6901 | -92.1084
2.6939 | -92.0969
2.6976 | -92.0257
2.7013 | -91.9052
2.705 | -91.7689
2.7088 | -91.6561
2.7125 | -91.5903
2.7162 | -91.5863
2.7199 | -91.6189
2.7237 | -91.6402
2.7274 | -91.6302
2.7311 | -91.6003
2.7349 | -91.5797
2.7386 | -91.5955
2.7423 | -91.6604
2.746 | -91.763
2.7498 | -91.8647
2.7535 | -91.9414
2.7572 | -91.9901
2.7609 | -92.0233
2.7647 | -92.046
2.7684 | -92.0532
2.7721 | -92.0501
2.7758 | -92.043
2.7796 | -92.0591
2.7833 | -92.1202
2.787 | -92.2314
2.7907 | -92.3697
2.7945 | -92.4835
2.7982 | -92.5563
2.8019 | -92.5733
2.8056 | -92.5275
2.8094 | -92.4126
2.8131 | -92.2616
2.8168 | -92.1059
2.8205 | -91.9492
2.8243 | -91.7871
2.828 | -91.6238
2.8317 | -91.4716
2.8355 | -91.3658
2.8392 | -91.3473
2.8429 | -91.4432
2.8466 | -91.6466
2.8504 | -91.9152
2.8541 | -92.1903
2.8578 | -92.4201
2.8615 | -92.5727
2.8653 | -92.6642
2.869 | -92.7414
2.8727 | -92.8045
2.8764 | -92.8244
2.8802 | -92.8035
2.8839 | -92.7613
2.8876 | -92.7221
2.8913 | -92.7232
2.8951 | -92.7755
2.8988 | -92.8726
2.9025 | -92.9837
2.9062 | -93.0772
2.91 | -93.1132
2.9137 | -93.0641
2.9174 | -92.9331
2.9211 | -92.7577
2.9249 | -92.5651
2.9286 | -92.382
2.9323 | -92.2362
2.9361 | -92.1581
2.9398 | -92.1453
2.9435 | -92.1782
2.9472 | -92.2365
2.951 | -92.301
2.9547 | -92.3513
2.9584 | -92.4008
2.9621 | -92.4594
2.9659 | -92.5201
2.9696 | -92.5942
2.9733 | -92.6994
2.977 | -92.8339
2.9808 | -92.9891
2.9845 | -93.1206
2.9882 | -93.1941
2.9919 | -93.1825
2.9957 | -93.0741
2.9994 | -92.9165
3.0031 | -92.7658
3.0068 | -92.6497
3.0106 | -92.5902
3.0143 | -92.5862
3.018 | -92.619
3.0217 | -92.6766
3.0255 | -92.767
3.0292 | -92.8881
3.0329 | -93.0396
3.0367 | -93.1865
3.0404 | -93.2957
3.0441 | -93.3602
3.0478 | -93.3702
3.0516 | -93.3494
3.0553 | -93.3165
3.059 | -93.2766
3.0627 | -93.2311
3.0665 | -93.1662
3.0702 | -93.07
3.0739 | -92.9483
3.0776 | -92.8385
3.0814 | -92.7533
3.0851 | -92.6919
3.0888 | -92.6309
3.0925 | -92.5478
3.0963 | -92.4683
3.1 | -92.4339
3.1037 | -92.4711
3.1074 | -92.5856
3.1112 | -92.7555
3.1149 | -92.9375
3.1186 | -93.1007
3.1224 | -93.2196
3.1261 | -93.2597
3.1298 | -93.2191
3.1335 | -93.122
3.1373 | -93.0093
3.141 | -92.925
3.1447 | -92.8963
3.1484 | -92.9481
3.1522 | -93.0819
3.1559 | -93.2696
3.1596 | -93.4766
3.1633 | -93.6687
3.1671 | -93.8082
3.1708 | -93.8903
3.1745 | -93.9501
3.1782 | -94.0106
3.182 | -94.0712
3.1857 | -94.1251
3.1894 | -94.1574
3.1931 | -94.1606
3.1969 | -94.1396
3.2006 | -94.1151
3.2043 | -94.0999
3.208 | -94.1079
3.2118 | -94.1609
3.2155 | -94.2705
3.2192 | -94.4204
3.223 | -94.5136
3.2267 | -94.5906
3.2304 | -94.6163
3.2341 | -94.6005
3.2379 | -94.56
3.2416 | -94.5092
3.2453 | -94.4568
3.249 | -94.4093
3.2528 | -94.3706
3.2565 | -94.3453
3.2602 | -94.3361
3.2639 | -94.3324
3.2677 | -94.3266
3.2714 | -94.308
3.2751 | -94.2653
3.2788 | -94.2043
3.2826 | -94.1283
3.2863 | -94.0608
3.29 | -94.022
3.2937 | -94.023
3.2975 | -94.0599
3.3012 | -94.1231
3.3049 | -94.2011
3.3086 | -94.2761
3.3124 | -94.3342
3.3161 | -94.3768
3.3198 | -94.4038
3.3236 | -94.4108
3.3273 | -94.4017
3.331 | -94.3795
3.3347 | -94.3467
3.3385 | -94.3161
3.3422 | -94.2894
3.3459 | -94.2801
3.3496 | -94.2976
3.3534 | -94.3388
3.3571 | -94.3909
3.3608 | -94.4423
3.3645 | -94.4883
3.3683 | -94.5247
3.372 | -94.5355
3.3757 | -94.5268
3.3794 | -94.4927
3.3832 | -94.4181
3.3869 | -94.3277
3.3906 | -94.2542
3.3943 | -94.2193
3.3981 | -94.2315
3.4018 | -94.286
3.4055 | -94.3717
3.4093 | -94.4801
3.413 | -94.5878
3.4167 | -94.669
3.4204 | -94.7062
3.4242 | -94.6936
3.4279 | -94.6559
3.4316 | -94.6143
3.4353 | -94.584
3.4391 | -94.5798
3.4428 | -94.5786
3.4465 | -94.5635
3.4502 | -94.5456
3.454 | -94.533
3.4577 | -94.5252
3.4614 | -94.5229
3.4651 | -94.5281
3.4689 | -94.5519
3.4726 | -94.5962
3.4763 | -94.6572
3.48 | -94.7242
3.4838 | -94.7799
3.4875 | -94.8139
3.4912 | -94.8314
3.4949 | -94.8411
3.4987 | -94.8632
3.5024 | -94.9164
3.5061 | -94.9813
3.5099 | -95.0493
3.5136 | -95.1174
3.5173 | -95.1679
3.521 | -95.1648
3.5248 | -95.1069
3.5285 | -94.9997
3.5322 | -94.8293
3.5359 | -94.6174
3.5397 | -94.4019
3.5434 | -94.2119
3.5471 | -94.0651
3.5508 | -93.9767
3.5546 | -93.9541
3.5583 | -93.99
3.562 | -94.0741
3.5657 | -94.1763
3.5695 | -94.2641
3.5732 | -94.3169
3.5769 | -94.3188
3.5806 | -94.2653
3.5844 | -94.1883
3.5881 | -94.1356
3.5918 | -94.1272
3.5955 | -94.1757
3.5993 | -94.2912
3.603 | -94.4573
3.6067 | -94.6463
3.6105 | -94.8359
3.6142 | -95.0074
3.6179 | -95.1394
3.6216 | -95.2084
3.6254 | -95.2115
3.6291 | -95.158
3.6328 | -95.0816
3.6365 | -95.0079
3.6403 | -94.9478
3.644 | -94.9076
3.6477 | -94.8858
3.6514 | -94.8713
3.6552 | -94.8585
3.6589 | -94.8491
3.6626 | -94.8393
3.6663 | -94.8212
3.6701 | -94.7968
3.6738 | -94.7784
3.6775 | -94.7725
3.6812 | -94.7695
3.685 | -94.7732
3.6887 | -94.7832
3.6924 | -94.7985
3.6961 | -94.8253
3.6999 | -94.8719
3.7036 | -94.9348
3.7073 | -95.0018
3.7111 | -95.0591
3.7148 | -95.1044
3.7185 | -95.1389
3.7222 | -95.1672
3.726 | -95.1897
3.7297 | -95.2085
3.7334 | -95.2299
3.7371 | -95.2547
3.7409 | -95.2773
3.7446 | -95.287
3.7483 | -95.2775
3.752 | -95.2496
3.7558 | -95.2038
3.7595 | -95.1623
3.7632 | -95.1301
3.7669 | -95.1149
3.7707 | -95.1226
3.7744 | -95.1478
3.7781 | -95.1836
3.7818 | -95.226
3.7856 | -95.2579
3.7893 | -95.2609
3.793 | -95.2284
3.7967 | -95.1701
3.8005 | -95.1007
3.8042 | -95.0379
3.8079 | -94.9966
3.8117 | -94.9895
3.8154 | -95.0282
3.8191 | -95.1128
3.8228 | -95.2368
3.8266 | -95.3807
3.8303 | -95.5208
3.834 | -95.6314
3.8377 | -95.7026
3.8415 | -95.7281
3.8452 | -95.7123
3.8489 | -95.6742
3.8526 | -95.629
3.8564 | -95.5822
3.8601 | -95.538
3.8638 | -95.5011
3.8675 | -95.4739
3.8713 | -95.4558
3.875 | -95.4429
3.8787 | -95.438
3.8824 | -95.4453
3.8862 | -95.4635
3.8899 | -95.5051
3.8936 | -95.5754
3.8973 | -95.6698
3.9011 | -95.7738
3.9048 | -95.8693
3.9085 | -95.9401
3.9123 | -95.9825
3.916 | -95.9901
3.9197 | -95.9685
3.9234 | -95.9306
3.9272 | -95.8842
3.9309 | -95.8295
3.9346 | -95.7665
3.9383 | -95.7003
3.9421 | -95.6382
3.9458 | -95.5813
3.9495 | -95.5294
3.9532 | -95.4785
3.957 | -95.4333
3.9607 | -95.3984
3.9644 | -95.3856
3.9681 | -95.3969
3.9719 | -95.4315
3.9756 | -95.4841
3.9793 | -95.5376
3.983 | -95.5746
3.9868 | -95.59
3.9905 | -95.5881
3.9942 | -95.5797
3.998 | -95.5678
4.0017 | -95.5562
4.0054 | -95.5487
4.0091 | -95.5531
4.0129 | -95.564
4.0166 | -95.5788
4.0203 | -95.6079
4.024 | -95.6526
4.0278 | -95.708
4.0315 | -95.7695
4.0352 | -95.8312
4.0389 | -95.884
4.0427 | -95.9297
4.0464 | -95.958
4.0501 | -95.961
4.0538 | -95.938
4.0576 | -95.9046
4.0613 | -95.8628
4.065 | -95.8059
4.0687 | -95.7373
4.0725 | -95.6616
4.0762 | -95.5872
4.0799 | -95.5165
4.0836 | -95.4497
4.0874 | -95.4004
4.0911 | -95.3638
4.0948 | -95.3379
4.0986 | -95.3163
4.1023 | -95.3064
4.106 | -95.312
4.1097 | -95.3266
4.1135 | -95.3516
4.1172 | -95.3907
4.1209 | -95.4357
4.1246 | -95.4941
4.1284 | -95.5706
4.1321 | -95.66
4.1358 | -95.7538
4.1395 | -95.8444
4.1433 | -95.9253
4.147 | -95.9935
4.1507 | -96.0438
4.1544 | -96.0778
4.1582 | -96.1107
4.1619 | -96.1385
4.1656 | -96.1571
4.1693 | -96.1705
4.1731 | -96.1936
4.1768 | -96.2211
4.1805 | -96.2493
4.1842 | -96.267
4.188 | -96.2689
4.1917 | -96.2522
4.1954 | -96.2239
4.1992 | -96.195
4.2029 | -96.176
4.2066 | -96.1778
4.2103 | -96.188
4.2141 | -96.2077
4.2178 | -96.2339
4.2215 | -96.2504
4.2252 | -96.2614
4.229 | -96.2714
4.2327 | -96.2851
4.2364 | -96.3066
4.2401 | -96.3426
4.2439 | -96.379
4.2476 | -96.4045
4.2513 | -96.4067
4.255 | -96.373
4.2588 | -96.2988
4.2625 | -96.1915
4.2662 | -96.0784
4.2699 | -95.9795
4.2737 | -95.9257
4.2774 | -95.9251
4.2811 | -95.9771
4.2848 | -96.0729
4.2886 | -96.2071
4.2923 | -96.3593
4.296 | -96.5111
4.2998 | -96.5965
4.3035 | -96.6891
4.3072 | -96.7472
4.3109 | -96.7807
4.3147 | -96.7909
4.3184 | -96.7764
4.3221 | -96.7499
4.3258 | -96.7124
4.3296 | -96.677
4.3333 | -96.6353
4.337 | -96.5938
4.3407 | -96.5544
4.3445 | -96.5183
4.3482 | -96.4877
4.3519 | -96.4681
4.3556 | -96.4553
4.3594 | -96.455
4.3631 | -96.4738
4.3668 | -96.515
4.3705 | -96.5898
4.3743 | -96.6912
4.378 | -96.8104
4.3817 | -96.9255
4.3854 | -97.0268
4.3892 | -97.1051
4.3929 | -97.1592
4.3966 | -97.1971
4.4004 | -97.2118
4.4041 | -97.2105
4.4078 | -97.2014
4.4115 | -97.1851
4.4153 | -97.1641
4.419 | -97.1488
4.4227 | -97.1372
4.4264 | -97.1267
4.4302 | -97.1144
4.4339 | -97.0926
4.4376 | -97.0616
4.4413 | -97.0203
4.4451 | -96.9677
4.4488 | -96.9074
4.4525 | -96.8505
4.4562 | -96.7981
4.46 | -96.7536
4.4637 | -96.7128
4.4674 | -96.6816
4.4711 | -96.6622
4.4749 | -96.6533
4.4786 | -96.6421
4.4823 | -96.6266
4.4861 | -96.6059
4.4898 | -96.57
4.4935 | -96.526
4.4972 | -96.4793
4.501 | -96.4279
4.5047 | -96.3735
4.5084 | -96.314
4.5121 | -96.2463
4.5159 | -96.1821
4.5196 | -96.1252
4.5233 | -96.0823
4.527 | -96.0585
4.5308 | -96.0629
4.5345 | -96.101
4.5382 | -96.1648
4.5419 | -96.2507
4.5457 | -96.3475
4.5494 | -96.4416
4.5531 | -96.525
4.5568 | -96.5875
4.5606 | -96.6273
4.5643 | -96.6383
4.568 | -96.6263
4.5717 | -96.612
4.5755 | -96.6076
4.5792 | -96.6183
4.5829 | -96.6469
4.5867 | -96.6892
4.5904 | -96.7453
4.5941 | -96.8104
4.5978 | -96.884
4.6016 | -96.9575
4.6053 | -97.0248
4.609 | -97.0962
4.6127 | -97.1729
4.6165 | -97.2611
4.6202 | -97.3592
4.6239 | -97.456
4.6276 | -97.5533
4.6314 | -97.649
4.6351 | -97.7395
4.6388 | -97.804
4.6425 | -97.8316
4.6463 | -97.827
4.65 | -97.8028
4.6537 | -97.7742
4.6574 | -97.7515
4.6612 | -97.731
4.6649 | -97.7191
4.6686 | -97.7136
4.6723 | -97.7126
4.6761 | -97.7026
4.6798 | -97.677
4.6835 | -97.6424
4.6873 | -97.6003
4.691 | -97.5674
4.6947 | -97.5483
4.6984 | -97.544
4.7022 | -97.5552
4.7059 | -97.5775
4.7096 | -97.6053
4.7133 | -97.6231
4.7171 | -97.6278
4.7208 | -97.6185
4.7245 | -97.596
4.7282 | -97.5703
4.732 | -97.5512
4.7357 | -97.5391
4.7394 | -97.5375
4.7431 | -97.5448
4.7469 | -97.5667
4.7506 | -97.5968
4.7543 | -97.6354
4.758 | -97.6762
4.7618 | -97.702
4.7655 | -97.7193
4.7692 | -97.7276
4.7729 | -97.7217
4.7767 | -97.698
4.7804 | -97.6642
4.7841 | -97.625
4.7879 | -97.5765
4.7916 | -97.5187
4.7953 | -97.4561
4.799 | -97.3885
4.8028 | -97.326
4.8065 | -97.2756
4.8102 | -97.2353
4.8139 | -97.2126
4.8177 | -97.2053
4.8214 | -97.2089
4.8251 | -97.2175
4.8288 | -97.2351
4.8326 | -97.2637
4.8363 | -97.3039
4.84 | -97.3318
4.8437 | -97.3563
4.8475 | -97.3833
4.8512 | -97.4071
4.8549 | -97.4161
4.8586 | -97.3988
4.8624 | -97.3569
4.8661 | -97.2981
4.8698 | -97.2258
4.8735 | -97.1461
4.8773 | -97.0682
4.881 | -97.0025
4.8847 | -96.9537
4.8885 | -96.9225
4.8922 | -96.9165
4.8959 | -96.9401
4.8996 | -96.991
4.9034 | -97.0561
4.9071 | -97.1359
4.9108 | -97.23
4.9145 | -97.3288
4.9183 | -97.4282
4.922 | -97.5222
4.9257 | -97.6068
4.9294 | -97.6813
4.9332 | -97.7424
4.9369 | -97.7931
4.9406 | -97.8297
4.9443 | -97.851
4.9481 | -97.862
4.9518 | -97.8649
4.9555 | -97.8522
4.9592 | -97.8163
4.963 | -97.7576
4.9667 | -97.6798
4.9704 | -97.5906
4.9742 | -97.5002
4.9779 | -97.4165
4.9816 | -97.3534
4.9853 | -97.3058
4.9891 | -97.2756
4.9928 | -97.2702
4.9965 | -97.2886
5.0002 | -97.3349
5.004 | -97.41
5.0077 | -97.5086
5.0114 | -97.6345
5.0151 | -97.7792
5.0189 | -97.9203
5.0226 | -98.0498
5.0263 | -98.1591
5.03 | -98.2381
5.0338 | -98.2778
5.0375 | -98.2837
5.0412 | -98.2622
5.0449 | -98.2036
5.0487 | -98.1139
5.0524 | -98.0059
5.0561 | -97.8902
5.0598 | -97.7707
5.0636 | -97.6508
5.0673 | -97.534
5.071 | -97.4308
5.0748 | -97.3417
5.0785 | -97.2618
5.0822 | -97.2007
5.0859 | -97.1636
5.0897 | -97.1356
5.0934 | -97.1282
5.0971 | -97.1523
5.1008 | -97.1951
5.1046 | -97.2518
5.1083 | -97.3293
5.112 | -97.4285
5.1157 | -97.5332
5.1195 | -97.6338
5.1232 | -97.7285
5.1269 | -97.807
5.1306 | -97.8595
5.1344 | -97.8839
5.1381 | -97.888
5.1418 | -97.8801
5.1455 | -97.858
5.1493 | -97.823
5.153 | -97.7851
5.1567 | -97.7556
5.1604 | -97.7366
5.1642 | -97.7295
5.1679 | -97.7439
5.1716 | -97.7681
5.1754 | -97.8026
5.1791 | -97.8501
5.1828 | -97.905
5.1865 | -97.9488
5.1903 | -97.9855
5.194 | -98.0188
5.1977 | -98.04
5.2014 | -98.0471
5.2052 | -98.0505
5.2089 | -98.0548
5.2126 | -98.0615
5.2163 | -98.0615
5.2201 | -98.059
5.2238 | -98.0676
5.2275 | -98.0829
5.2312 | -98.0983
5.235 | -98.1058
5.2387 | -98.1103
5.2424 | -98.1231
5.2461 | -98.1393
5.2499 | -98.1545
5.2536 | -98.1668
5.2573 | -98.1772
5.2611 | -98.1871
5.2648 | -98.1955
5.2685 | -98.1902
5.2722 | -98.1778
5.276 | -98.1621
5.2797 | -98.1421
5.2834 | -98.1177
5.2871 | -98.0957
5.2909 | -98.0852
5.2946 | -98.0876
5.2983 | -98.093
5.302 | -98.1042
5.3058 | -98.1205
5.3095 | -98.1392
5.3132 | -98.1461
5.3169 | -98.1266
5.3207 | -98.0742
5.3244 | -97.9992
5.3281 | -97.9052
5.3318 | -97.7971
5.3356 | -97.6803
5.3393 | -97.5627
5.343 | -97.4573
5.3467 | -97.3721
5.3505 | -97.3075
5.3542 | -97.2685
5.3579 | -97.2528
5.3617 | -97.2558
5.3654 | -97.2772
5.3691 | -97.3119
5.3728 | -97.3866
5.3766 | -97.4475
5.3803 | -97.5265
5.384 | -97.609
5.3877 | -97.6928
5.3915 | -97.7787
5.3952 | -97.8672
5.3989 | -97.9484
5.4026 | -98.0148
5.4064 | -98.0571
5.4101 | -98.0793
5.4138 | -98.0788
5.4175 | -98.0583
5.4213 | -98.0262
5.425 | -97.9883
5.4287 | -97.9509
5.4324 | -97.926
5.4362 | -97.9146
5.4399 | -97.9254
5.4436 | -97.9583
5.4473 | -98.0077
5.4511 | -98.0667
5.4548 | -98.1279
5.4585 | -98.1828
5.4623 | -98.2292
5.466 | -98.2662
5.4697 | -98.2919
5.4734 | -98.3078
5.4772 | -98.313
5.4809 | -98.3102
5.4846 | -98.3041
5.4883 | -98.308
5.4921 | -98.3225
5.4958 | -98.344
5.4995 | -98.3761
5.5032 | -98.416
5.507 | -98.46
5.5107 | -98.5085
5.5144 | -98.557
5.5181 | -98.6007
5.5219 | -98.6378
5.5256 | -98.6626
5.5293 | -98.6779
5.533 | -98.6796
5.5368 | -98.6659
5.5405 | -98.6347
5.5442 | -98.5917
5.5479 | -98.5314
5.5517 | -98.4535
5.5554 | -98.3682
5.5591 | -98.2802
5.5629 | -98.1963
5.5666 | -98.1222
5.5703 | -98.0612
5.574 | -98.0076
5.5778 | -97.9675
5.5815 | -97.9516
5.5852 | -97.9549
5.5889 | -97.9723
5.5927 | -98.0027
5.5964 | -98.0358
5.6001 | -98.0667
5.6038 | -98.0938
5.6076 | -98.1159
5.6113 | -98.1342
5.615 | -98.1446
5.6187 | -98.1525
5.6225 | -98.1588
5.6262 | -98.1588
5.6299 | -98.1586
5.6336 | -98.1554
5.6374 | -98.1457
5.6411 | -98.1293
5.6448 | -98.1072
5.6485 | -98.0802
5.6523 | -98.0502
5.656 | -98.0212
5.6597 | -97.9934
5.6635 | -97.9731
5.6672 | -97.9576
5.6709 | -97.9473
5.6746 | -97.9425
5.6784 | -97.9456
5.6821 | -97.9661
5.6858 | -97.9971
5.6895 | -98.0296
5.6933 | -98.054
5.697 | -98.065
5.7007 | -98.0714
5.7044 | -98.0719
5.7082 | -98.0699
5.7119 | -98.0659
5.7156 | -98.0529
5.7193 | -98.0428
5.7231 | -98.0437
5.7268 | -98.0571
5.7305 | -98.0786
5.7342 | -98.1055
5.738 | -98.133
5.7417 | -98.1566
5.7454 | -98.1752
5.7492 | -98.1815
5.7529 | -98.1748
5.7566 | -98.1704
5.7603 | -98.1684
5.7641 | -98.1601
5.7678 | -98.1458
5.7715 | -98.1316
5.7752 | -98.1195
5.779 | -98.1116
5.7827 | -98.106
5.7864 | -98.1015
5.7901 | -98.1051
5.7939 | -98.1139
5.7976 | -98.1235
5.8013 | -98.1382
5.805 | -98.1559
5.8088 | -98.1657
5.8125 | -98.1627
5.8162 | -98.1402
5.8199 | -98.0925
5.8237 | -98.0127
5.8274 | -97.9142
5.8311 | -97.797
5.8348 | -97.6677
5.8386 | -97.533
5.8423 | -97.3926
5.846 | -97.257
5.8498 | -97.134
5.8535 | -97.03
5.8572 | -96.9555
5.8609 | -96.9098
5.8647 | -96.8842
5.8684 | -96.8665
5.8721 | -96.8563
5.8758 | -96.8541
5.8796 | -96.852
5.8833 | -96.8581
5.887 | -96.866
5.8907 | -96.8742
5.8945 | -96.8918
5.8982 | -96.9226
5.9019 | -96.9734
5.9056 | -97.041
5.9094 | -97.1433
5.9131 | -97.24
5.9168 | -97.3424
5.9205 | -97.4493
5.9243 | -97.5583
5.928 | -97.6681
5.9317 | -97.7787
5.9354 | -97.891
5.9392 | -97.9994
5.9429 | -98.0915
5.9466 | -98.1623
5.9504 | -98.215
5.9541 | -98.2604
5.9578 | -98.3049
5.9615 | -98.345
5.9653 | -98.3815
5.969 | -98.4277
5.9727 | -98.4757
5.9764 | -98.5319
5.9802 | -98.6013
5.9839 | -98.6782
5.9876 | -98.7572
5.9913 | -98.8347
5.9951 | -98.9136
5.9988 | -98.9955
6.0025 | -99.072
6.0062 | -99.1409
6.01 | -99.2043
6.0137 | -99.2707
6.0174 | -99.3299
6.0211 | -99.3774
6.0249 | -99.4161
6.0286 | -99.4394
6.0323 | -99.4464
6.036 | -99.4438
6.0398 | -99.4253
6.0435 | -99.3942
6.0472 | -99.3587
6.051 | -99.3224
6.0547 | -99.2855
6.0584 | -99.2644
6.0621 | -99.2614
6.0659 | -99.2761
6.0696 | -99.3089
6.0733 | -99.3582
6.077 | -99.4183
6.0808 | -99.4837
6.0845 | -99.5431
6.0882 | -99.5918
6.0919 | -99.619
6.0957 | -99.6224
6.0994 | -99.6094
6.1031 | -99.5839
6.1068 | -99.5458
6.1106 | -99.505
6.1143 | -99.4672
6.118 | -99.4348
6.1217 | -99.4108
6.1255 | -99.3998
6.1292 | -99.406
6.1329 | -99.4304
6.1366 | -99.4737
6.1404 | -99.5273
6.1441 | -99.586
6.1478 | -99.6509
6.1516 | -99.713
6.1553 | -99.7759
6.159 | -99.8326
6.1627 | -99.8771
6.1665 | -99.9062
6.1702 | -99.9146
6.1739 | -99.8985
6.1776 | -99.8669
6.1814 | -99.8153
6.1851 | -99.7437
6.1888 | -99.6578
6.1925 | -99.5534
6.1963 | -99.4388
6.2 | -99.3214
6.2037 | -99.2135
6.2074 | -99.1197
6.2112 | -99.0472
6.2149 | -99.0018
6.2186 | -98.9808
6.2223 | -98.9793
6.2261 | -98.9975
6.2298 | -99.0266
6.2335 | -99.0661
6.2373 | -99.1117
6.241 | -99.154
6.2447 | -99.192
6.2484 | -99.222
6.2522 | -99.2425
6.2559 | -99.2594
6.2596 | -99.2679
6.2633 | -99.2762
6.2671 | -99.2893
6.2708 | -99.2963
6.2745 | -99.3069
6.2782 | -99.3238
6.282 | -99.3498
6.2857 | -99.3891
6.2894 | -99.44
6.2931 | -99.4931
6.2969 | -99.5395
6.3006 | -99.5804
6.3043 | -99.6131
6.308 | -99.6373
6.3118 | -99.6494
6.3155 | -99.6491
6.3192 | -99.6383
6.3229 | -99.6143
6.3267 | -99.5743
6.3304 | -99.5264
6.3341 | -99.4811
6.3379 | -99.4375
6.3416 | -99.402
6.3453 | -99.376
6.349 | -99.3582
6.3528 | -99.3614
6.3565 | -99.3828
6.3602 | -99.4188
6.3639 | -99.4735
6.3677 | -99.5442
6.3714 | -99.6294
6.3751 | -99.7158
6.3788 | -99.8054
6.3826 | -99.8984
6.3863 | -99.9957
6.39 | -100.0916
6.3937 | -100.1717
6.3975 | -100.2342
6.4012 | -100.279
6.4049 | -100.3103
6.4086 | -100.3228
6.4124 | -100.3229
6.4161 | -100.3186
6.4198 | -100.3099
6.4235 | -100.3007
6.4273 | -100.2912
6.431 | -100.2858
6.4347 | -100.2814
6.4385 | -100.2845
6.4422 | -100.2905
6.4459 | -100.2985
6.4496 | -100.3119
6.4534 | -100.3205
6.4571 | -100.327
6.4608 | -100.3294
6.4645 | -100.3292
6.4683 | -100.3235
6.472 | -100.3154
6.4757 | -100.3026
6.4794 | -100.28
6.4832 | -100.2542
6.4869 | -100.2304
6.4906 | -100.212
6.4943 | -100.1915
6.4981 | -100.1735
6.5018 | -100.1574
6.5055 | -100.1436
6.5092 | -100.1265
6.513 | -100.11
6.5167 | -100.1008
6.5204 | -100.0942
6.5241 | -100.0933
6.5279 | -100.1005
6.5316 | -100.1141
6.5353 | -100.1303
6.5391 | -100.1426
6.5428 | -100.1518
6.5465 | -100.1563
6.5502 | -100.1556
6.554 | -100.1524
6.5577 | -100.1445
6.5614 | -100.1354
6.5651 | -100.1246
6.5689 | -100.1114
6.5726 | -100.0908
6.5763 | -100.0632
6.58 | -100.0349
6.5838 | -100.0065
6.5875 | -99.9784
6.5912 | -99.9557
6.5949 | -99.9403
6.5987 | -99.9308
6.6024 | -99.9243
6.6061 | -99.9255
6.6098 | -99.9384
6.6136 | -99.956
6.6173 | -99.9706
6.621 | -99.9862
6.6247 | -100.0062
6.6285 | -100.0262
6.6322 | -100.0549
6.6359 | -100.086
6.6397 | -100.1231
6.6434 | -100.1579
6.6471 | -100.195
6.6508 | -100.2356
6.6546 | -100.2722
6.6583 | -100.3081
6.662 | -100.3427
6.6657 | -100.3743
6.6695 | -100.4069
6.6732 | -100.4413
6.6769 | -100.4696
6.6806 | -100.4881
6.6844 | -100.4984
6.6881 | -100.5008
6.6918 | -100.4973
6.6955 | -100.4896
6.6993 | -100.4733
6.703 | -100.4522
6.7067 | -100.4246
6.7104 | -100.3945
6.7142 | -100.3645
6.7179 | -100.3402
6.7216 | -100.3217
6.7253 | -100.3078
6.7291 | -100.3051
6.7328 | -100.3075
6.7365 | -100.3239
6.7403 | -100.3464
6.744 | -100.3782
6.7477 | -100.4183
6.7514 | -100.4588
6.7552 | -100.4957
6.7589 | -100.5334
6.7626 | -100.5644
6.7663 | -100.5805
6.7701 | -100.5842
6.7738 | -100.5761
6.7775 | -100.5597
6.7812 | -100.535
6.785 | -100.4976
6.7887 | -100.4521
6.7924 | -100.3996
6.7961 | -100.3451
6.7999 | -100.2913
6.8036 | -100.2467
6.8073 | -100.2059
6.811 | -100.1696
6.8148 | -100.1414
6.8185 | -100.1192
6.8222 | -100.1006
6.826 | -100.0868
6.8297 | -100.0798
6.8334 | -100.0787
6.8371 | -100.0775
6.8409 | -100.083
6.8446 | -100.0934
6.8483 | -100.1074
6.852 | -100.1298
6.8558 | -100.1628
6.8595 | -100.1976
6.8632 | -100.236
6.8669 | -100.2847
6.8707 | -100.3402
6.8744 | -100.3964
6.8781 | -100.4499
6.8818 | -100.4984
6.8856 | -100.5415
6.8893 | -100.5806
6.893 | -100.6119
6.8967 | -100.6413
6.9005 | -100.6641
6.9042 | -100.6676
6.9079 | -100.662
6.9116 | -100.6487
6.9154 | -100.6243
6.9191 | -100.59
6.9228 | -100.5442
6.9266 | -100.4948
6.9303 | -100.4432
6.934 | -100.387
6.9377 | -100.3287
6.9415 | -100.2701
6.9452 | -100.2116
6.9489 | -100.1581
6.9526 | -100.1113
6.9564 | -100.0701
6.9601 | -100.0397
6.9638 | -100.019
6.9675 | -100.0089
6.9713 | -100.0078
6.975 | -100.0223
6.9787 | -100.0469
6.9824 | -100.0812
6.9862 | -100.1425
6.9899 | -100.1867
6.9936 | -100.2308
6.9973 | -100.2764
7.0011 | -100.3188
7.0048 | -100.3614
7.0085 | -100.4034
7.0122 | -100.4351
7.016 | -100.4564
7.0197 | -100.4708
7.0234 | -100.4708
7.0272 | -100.4648
7.0309 | -100.4457
7.0346 | -100.4236
7.0383 | -100.4068
7.0421 | -100.3928
7.0458 | -100.3822
7.0495 | -100.3742
7.0532 | -100.367
7.057 | -100.3692
7.0607 | -100.3826
7.0644 | -100.3985
7.0681 | -100.4205
7.0719 | -100.4428
7.0756 | -100.4685
7.0793 | -100.4945
7.083 | -100.5239
7.0868 | -100.5532
7.0905 | -100.5808
7.0942 | -100.6104
7.0979 | -100.6395
7.1017 | -100.6646
7.1054 | -100.6828
7.1091 | -100.6848
7.1128 | -100.6782
7.1166 | -100.6632
7.1203 | -100.6441
7.124 | -100.6145
7.1278 | -100.5761
7.1315 | -100.5384
7.1352 | -100.4995
7.1389 | -100.4619
7.1427 | -100.422
7.1464 | -100.3798
7.1501 | -100.3391
7.1538 | -100.3004
7.1576 | -100.2695
7.1613 | -100.2458
7.165 | -100.225
7.1687 | -100.2101
7.1725 | -100.2082
7.1762 | -100.2138
7.1799 | -100.2303
7.1836 | -100.2536
7.1874 | -100.2844
7.1911 | -100.3229
7.1948 | -100.3666
7.1985 | -100.4064
7.2023 | -100.4469
7.206 | -100.4904
7.2097 | -100.5327
7.2135 | -100.5728
7.2172 | -100.6133
7.2209 | -100.6524
7.2246 | -100.6894
7.2284 | -100.7177
7.2321 | -100.7472
7.2358 | -100.7773
7.2395 | -100.801
7.2433 | -100.8217
7.247 | -100.8452
7.2507 | -100.8685
7.2544 | -100.8982
7.2582 | -100.9317
7.2619 | -100.9608
7.2656 | -100.9896
7.2693 | -101.0176
7.2731 | -101.0456
7.2768 | -101.074
7.2805 | -101.1067
7.2842 | -101.1458
7.288 | -101.182
7.2917 | -101.2082
7.2954 | -101.2297
7.2991 | -101.247
7.3029 | -101.2614
7.3066 | -101.2753
7.3103 | -101.2792
7.3141 | -101.2727
7.3178 | -101.2586
7.3215 | -101.2413
7.3252 | -101.2205
7.329 | -101.1928
7.3327 | -101.1645
7.3364 | -101.1372
7.3401 | -101.108
7.3439 | -101.0815
7.3476 | -101.0574
7.3513 | -101.0365
7.355 | -101.0163
7.3588 | -101.0059
7.3625 | -100.9957
7.3662 | -100.9914
7.3699 | -100.9997
7.3737 | -101.0088
7.3774 | -101.0172
7.3811 | -101.0273
7.3848 | -101.0407
7.3886 | -101.0526
7.3923 | -101.0665
7.396 | -101.0852
7.3997 | -101.1108
7.4035 | -101.1309
7.4072 | -101.147
7.4109 | -101.1651
7.4147 | -101.1788
7.4184 | -101.1928
7.4221 | -101.2094
7.4258 | -101.2234
7.4296 | -101.2345
7.4333 | -101.2457
7.437 | -101.2538
7.4407 | -101.2628
7.4445 | -101.2804
7.4482 | -101.2986
7.4519 | -101.319
7.4556 | -101.3496
7.4594 | -101.3873
7.4631 | -101.4279
7.4668 | -101.4699
7.4705 | -101.5228
7.4743 | -101.5722
7.478 | -101.6116
7.4817 | -101.6506
7.4854 | -101.6849
7.4892 | -101.7097
7.4929 | -101.7296
7.4966 | -101.7436
7.5003 | -101.7493
7.5041 | -101.7497
7.5078 | -101.7499
7.5115 | -101.7565
7.5153 | -101.7673
7.519 | -101.7706
7.5227 | -101.7676
7.5264 | -101.7662
7.5302 | -101.7637
7.5339 | -101.7581
7.5376 | -101.7467
7.5413 | -101.7317
7.5451 | -101.7087
7.5488 | -101.6846
7.5525 | -101.6582
7.5562 | -101.6233
7.56 | -101.5888
7.5637 | -101.5561
7.5674 | -101.522
7.5711 | -101.4869
7.5749 | -101.4582
7.5786 | -101.4333
7.5823 | -101.4094
7.586 | -101.3914
7.5898 | -101.3798
7.5935 | -101.3679
7.5972 | -101.3594
7.601 | -101.3538
7.6047 | -101.3507
7.6084 | -101.356
7.6121 | -101.3674
7.6159 | -101.3858
7.6196 | -101.4105
7.6233 | -101.4396
7.627 | -101.4749
7.6308 | -101.517
7.6345 | -101.5664
7.6382 | -101.6203
7.6419 | -101.6777
7.6457 | -101.7305
7.6494 | -101.7795
7.6531 | -101.826
7.6568 | -101.8628
7.6606 | -101.8942
7.6643 | -101.92
7.668 | -101.9369
7.6717 | -101.9469
7.6755 | -101.9436
7.6792 | -101.9284
7.6829 | -101.9092
7.6866 | -101.8863
7.6904 | -101.8552
7.6941 | -101.8163
7.6978 | -101.7731
7.7016 | -101.7281
7.7053 | -101.6831
7.709 | -101.6424
7.7127 | -101.6017
7.7165 | -101.5683
7.7202 | -101.5381
7.7239 | -101.516
7.7276 | -101.5079
7.7314 | -101.509
7.7351 | -101.5173
7.7388 | -101.5312
7.7425 | -101.5527
7.7463 | -101.5787
7.75 | -101.5985
7.7537 | -101.6114
7.7574 | -101.6226
7.7612 | -101.632
7.7649 | -101.631
7.7686 | -101.6223
7.7723 | -101.6058
7.7761 | -101.5813
7.7798 | -101.5461
7.7835 | -101.5053
7.7872 | -101.4584
7.791 | -101.417
7.7947 | -101.3752
7.7984 | -101.3312
7.8022 | -101.2907
7.8059 | -101.2575
7.8096 | -101.2338
7.8133 | -101.2217
7.8171 | -101.2187
7.8208 | -101.2208
7.8245 | -101.2272
7.8282 | -101.2458
7.832 | -101.2685
7.8357 | -101.2966
7.8394 | -101.3295
7.8431 | -101.3627
7.8469 | -101.3936
7.8506 | -101.4227
7.8543 | -101.4507
7.858 | -101.4736
7.8618 | -101.4918
7.8655 | -101.5028
7.8692 | -101.4979
7.8729 | -101.4736
7.8767 | -101.4404
7.8804 | -101.4053
7.8841 | -101.3663
7.8878 | -101.325
7.8916 | -101.2796
7.8953 | -101.2296
7.899 | -101.1816
7.9028 | -101.1398
7.9065 | -101.0987
7.9102 | -101.0647
7.9139 | -101.0387
7.9177 | -101.0212
7.9214 | -101.0156
7.9251 | -101.0202
7.9288 | -101.0358
7.9326 | -101.0613
7.9363 | -101.1
7.94 | -101.1542
7.9437 | -101.212
7.9475 | -101.2757
7.9512 | -101.3436
7.9549 | -101.4062
7.9586 | -101.4684
7.9624 | -101.5264
7.9661 | -101.576
7.9698 | -101.6185
7.9735 | -101.6434
7.9773 | -101.6595
7.981 | -101.6713
7.9847 | -101.6711
7.9885 | -101.665
7.9922 | -101.6499
7.9959 | -101.6253
7.9996 | -101.6012
8.0034 | -101.5838
8.0071 | -101.5767
8.0108 | -101.5851
8.0145 | -101.6017
8.0183 | -101.6196
8.022 | -101.6415
8.0257 | -101.6729
8.0294 | -101.709
8.0332 | -101.74
8.0369 | -101.7716
8.0406 | -101.8066
8.0443 | -101.8383
8.0481 | -101.8692
8.0518 | -101.8942
8.0555 | -101.9078
8.0592 | -101.9198
8.063 | -101.9078
8.0667 | -101.9039
8.0704 | -101.8951
8.0741 | -101.8871
8.0779 | -101.8705
8.0816 | -101.8447
8.0853 | -101.8099
8.0891 | -101.7704
8.0928 | -101.7376
8.0965 | -101.6992
8.1002 | -101.6565
8.104 | -101.6161
8.1077 | -101.5785
8.1114 | -101.5426
8.1151 | -101.5093
8.1189 | -101.4717
8.1226 | -101.4368
8.1263 | -101.4089
8.13 | -101.3853
8.1338 | -101.3692
8.1375 | -101.3638
8.1412 | -101.3661
8.1449 | -101.3725
8.1487 | -101.3842
8.1524 | -101.4073
8.1561 | -101.4287
8.1598 | -101.4488
8.1636 | -101.4738
8.1673 | -101.5028
8.171 | -101.5327
8.1747 | -101.5635
8.1785 | -101.5947
8.1822 | -101.627
8.1859 | -101.6619
8.1897 | -101.6995
8.1934 | -101.7411
8.1971 | -101.7796
8.2008 | -101.8207
8.2046 | -101.8679
8.2083 | -101.9154
8.212 | -101.9563
8.2157 | -101.9919
8.2195 | -102.0293
8.2232 | -102.0669
8.2269 | -102.1071
8.2306 | -102.1419
8.2344 | -102.1729
8.2381 | -102.2017
8.2418 | -102.236
8.2455 | -102.2649
8.2493 | -102.2882
8.253 | -102.3127
8.2567 | -102.3391
8.2604 | -102.3606
8.2642 | -102.3733
8.2679 | -102.3773
8.2716 | -102.3799
8.2753 | -102.3769
8.2791 | -102.3707
8.2828 | -102.3623
8.2865 | -102.3473
8.2903 | -102.3299
8.294 | -102.3114
8.2977 | -102.2966
8.3014 | -102.2871
8.3052 | -102.2812
8.3089 | -102.2746
8.3126 | -102.2696
8.3163 | -102.2673
8.3201 | -102.2657
8.3238 | -102.2611
8.3275 | -102.2503
8.3312 | -102.2388
8.335 | -102.2235
8.3387 | -102.2053
8.3424 | -102.1859
8.3461 | -102.1627
8.3499 | -102.138
8.3536 | -102.1137
8.3573 | -102.0895
8.361 | -102.0648
8.3648 | -102.0367
8.3685 | -102.0089
8.3722 | -101.983
8.3759 | -101.9571
8.3797 | -101.9316
8.3834 | -101.9085
8.3871 | -101.89
8.3909 | -101.8735
8.3946 | -101.8546
8.3983 | -101.8379
8.402 | -101.8208
8.4058 | -101.8028
8.4095 | -101.7889
8.4132 | -101.7737
8.4169 | -101.7599
8.4207 | -101.742
8.4244 | -101.7197
8.4281 | -101.6968
8.4318 | -101.6768
8.4356 | -101.6588
8.4393 | -101.6399
8.443 | -101.6264
8.4467 | -101.6183
8.4505 | -101.6213
8.4542 | -101.6234
8.4579 | -101.6351
8.4616 | -101.6545
8.4654 | -101.6764
8.4691 | -101.7049
8.4728 | -101.7397
8.4765 | -101.7809
8.4803 | -101.8239
8.484 | -101.8673
8.4877 | -101.9103
8.4915 | -101.9525
8.4952 | -101.9883
8.4989 | -102.0212
8.5026 | -102.0505
8.5064 | -102.0763
8.5101 | -102.0932
8.5138 | -102.1052
8.5175 | -102.1161
8.5213 | -102.1221
8.525 | -102.1235
8.5287 | -102.1169
8.5324 | -102.1027
8.5362 | -102.0832
8.5399 | -102.0576
8.5436 | -102.0289
8.5473 | -101.9933
8.5511 | -101.9539
8.5548 | -101.9141
8.5585 | -101.8763
8.5622 | -101.8466
8.566 | -101.8266
8.5697 | -101.8134
8.5734 | -101.8017
8.5772 | -101.8008
8.5809 | -101.8085
8.5846 | -101.8227
8.5883 | -101.8433
8.5921 | -101.8669
8.5958 | -101.8921
8.5995 | -101.9048
8.6032 | -101.9156
8.607 | -101.9244
8.6107 | -101.9293
8.6144 | -101.9287
8.6181 | -101.9246
8.6219 | -101.9197
8.6256 | -101.91
8.6293 | -101.8981
8.633 | -101.8902
8.6368 | -101.8867
8.6405 | -101.8828
8.6442 | -101.8764
8.6479 | -101.8732
8.6517 | -101.8767
8.6554 | -101.8858
8.6591 | -101.8979
8.6628 | -101.9098
8.6666 | -101.9171
8.6703 | -101.9221
8.674 | -101.9307
8.6778 | -101.9345
8.6815 | -101.9297
8.6852 | -101.9198
8.6889 | -101.9014
8.6927 | -101.8765
8.6964 | -101.8414
8.7001 | -101.795
8.7038 | -101.7458
8.7076 | -101.6895
8.7113 | -101.6302
8.715 | -101.5678
8.7187 | -101.5031
8.7225 | -101.4408
8.7262 | -101.3804
8.7299 | -101.3243
8.7336 | -101.2692
8.7374 | -101.2156
8.7411 | -101.1699
8.7448 | -101.1319
8.7485 | -101.1003
8.7523 | -101.0742
8.756 | -101.056
8.7597 | -101.0468
8.7634 | -101.0426
8.7672 | -101.0452
8.7709 | -101.0585
8.7746 | -101.0812
8.7784 | -101.1109
8.7821 | -101.1502
8.7858 | -101.1972
8.7895 | -101.248
8.7933 | -101.3051
8.797 | -101.3719
8.8007 | -101.4392
8.8044 | -101.5032
8.8082 | -101.5673
8.8119 | -101.6275
8.8156 | -101.6802
8.8193 | -101.731
8.8231 | -101.7777
8.8268 | -101.8212
8.8305 | -101.8599
8.8342 | -101.8876
8.838 | -101.9053
8.8417 | -101.9172
8.8454 | -101.9288
8.8491 | -101.9372
8.8529 | -101.9366
8.8566 | -101.9371
8.8603 | -101.9397
8.864 | -101.9405
8.8678 | -101.9439
8.8715 | -101.9429
8.8752 | -101.933
8.879 | -101.9263
8.8827 | -101.9287
8.8864 | -101.9332
8.8901 | -101.9382
8.8939 | -101.9435
8.8976 | -101.9528
8.9013 | -101.9619
8.905 | -101.973
8.9088 | -101.9837
8.9125 | -101.9887
8.9162 | -101.9904
8.9199 | -101.9933
8.9237 | -101.9959
8.9274 | -101.9952
8.9311 | -101.9931
8.9348 | -101.9866
8.9386 | -101.9797
8.9423 | -101.9676
8.946 | -101.9581
8.9497 | -101.9468
8.9535 | -101.9353
8.9572 | -101.9289
8.9609 | -101.9256
8.9647 | -101.9251
8.9684 | -101.9302
8.9721 | -101.9364
8.9758 | -101.943
8.9796 | -101.9551
8.9833 | -101.9761
8.987 | -102.0028
8.9907 | -102.0289
8.9945 | -102.0631
8.9982 | -102.1007
9.0019 | -102.1416
9.0056 | -102.1911
9.0094 | -102.2351
9.0131 | -102.2709
9.0168 | -102.3078
9.0205 | -102.3436
9.0243 | -102.375
9.028 | -102.398
9.0317 | -102.4151
9.0354 | -102.4252
9.0392 | -102.4311
9.0429 | -102.4361
9.0466 | -102.4371
9.0503 | -102.4338
9.0541 | -102.4256
9.0578 | -102.4132
9.0615 | -102.4044
9.0653 | -102.4019
9.069 | -102.3975
9.0727 | -102.4006
9.0764 | -102.4045
9.0802 | -102.4115
9.0839 | -102.4215
9.0876 | -102.4334
9.0913 | -102.4477
9.0951 | -102.4611
9.0988 | -102.4791
9.1025 | -102.5011
9.1062 | -102.5217
9.11 | -102.5377
9.1137 | -102.5433
9.1174 | -102.5405
9.1211 | -102.5346
9.1249 | -102.5194
9.1286 | -102.4964
9.1323 | -102.4688
9.136 | -102.4345
9.1398 | -102.4101
9.1435 | -102.3886
9.1472 | -102.3747
9.1509 | -102.3622
9.1547 | -102.3548
9.1584 | -102.3533
9.1621 | -102.359
9.1659 | -102.3716
9.1696 | -102.3925
9.1733 | -102.4197
9.177 | -102.4485
9.1808 | -102.4831
9.1845 | -102.5224
9.1882 | -102.5633
9.1919 | -102.5993
9.1957 | -102.6219
9.1994 | -102.638
9.2031 | -102.6547
9.2068 | -102.6657
9.2106 | -102.6677
9.2143 | -102.6634
9.218 | -102.657
9.2217 | -102.6425
9.2255 | -102.6202
9.2292 | -102.5953
9.2329 | -102.5644
9.2366 | -102.5318
9.2404 | -102.4994
9.2441 | -102.4693
9.2478 | -102.4451
9.2515 | -102.4225
9.2553 | -102.4026
9.259 | -102.39
9.2627 | -102.382
9.2665 | -102.3796
9.2702 | -102.3852
9.2739 | -102.3919
9.2776 | -102.3992
9.2814 | -102.4107
9.2851 | -102.4201
9.2888 | -102.4312
9.2925 | -102.4447
9.2963 | -102.4585
9.3 | -102.4675
9.3037 | -102.4757
9.3074 | -102.4834
9.3112 | -102.4904
9.3149 | -102.4922
9.3186 | -102.4897
9.3223 | -102.4891
9.3261 | -102.4891
9.3298 | -102.4935
9.3335 | -102.5034
9.3372 | -102.5075
9.341 | -102.5082
9.3447 | -102.5085
9.3484 | -102.5128
9.3521 | -102.5169
9.3559 | -102.5206
9.3596 | -102.5291
9.3633 | -102.5363
9.3671 | -102.5402
9.3708 | -102.5445
9.3745 | -102.5509
9.3782 | -102.5596
9.382 | -102.5683
9.3857 | -102.5752
9.3894 | -102.5815
9.3931 | -102.5852
9.3969 | -102.5853
9.4006 | -102.5904
9.4043 | -102.5917
9.408 | -102.5915
9.4118 | -102.5955
9.4155 | -102.5977
9.4192 | -102.5987
9.4229 | -102.6
9.4267 | -102.6001
9.4304 | -102.5988
9.4341 | -102.5988
9.4378 | -102.5972
9.4416 | -102.5956
9.4453 | -102.5943
9.449 | -102.592
9.4527 | -102.5966
9.4565 | -102.5986
9.4602 | -102.5939
9.4639 | -102.5909
9.4677 | -102.5875
9.4714 | -102.5831
9.4751 | -102.5827
9.4788 | -102.5801
9.4826 | -102.5772
9.4863 | -102.573
9.49 | -102.5696
9.4937 | -102.5662
9.4975 | -102.5657
9.5012 | -102.5713
9.5049 | -102.5762
9.5086 | -102.5807
9.5124 | -102.5845
9.5161 | -102.5897
9.5198 | -102.5923
9.5235 | -102.5958
9.5273 | -102.6017
9.531 | -102.6077
9.5347 | -102.6123
9.5384 | -102.6073
9.5422 | -102.603
9.5459 | -102.5958
9.5496 | -102.5815
9.5534 | -102.5633
9.5571 | -102.5407
9.5608 | -102.5157
9.5645 | -102.4873
9.5683 | -102.4544
9.572 | -102.4211
9.5757 | -102.3895
9.5794 | -102.3588
9.5832 | -102.3286
9.5869 | -102.3041
9.5906 | -102.2841
9.5943 | -102.2682
9.5981 | -102.2604
9.6018 | -102.2586
9.6055 | -102.2672
9.6092 | -102.2863
9.613 | -102.3087
9.6167 | -102.3384
9.6204 | -102.3753
9.6241 | -102.417
9.6279 | -102.4666
9.6316 | -102.5198
9.6353 | -102.5715
9.639 | -102.6301
9.6428 | -102.6926
9.6465 | -102.7547
9.6502 | -102.8132
9.654 | -102.8677
9.6577 | -102.9224
9.6614 | -102.9688
9.6651 | -103.0004
9.6689 | -103.021
9.6726 | -103.014
9.6763 | -103.0247
9.68 | -103.0297
9.6838 | -103.0289
9.6875 | -103.0204
9.6912 | -103.0058
9.6949 | -102.9856
9.6987 | -102.9616
9.7024 | -102.9347
9.7061 | -102.9039
9.7098 | -102.8764
9.7136 | -102.8467
9.7173 | -102.8203
9.721 | -102.795
9.7247 | -102.7733
9.7285 | -102.758
9.7322 | -102.7479
9.7359 | -102.7487
9.7396 | -102.7558
9.7434 | -102.7735
9.7471 | -102.7978
9.7508 | -102.8298
9.7546 | -102.8703
9.7583 | -102.9172
9.762 | -102.9746
9.7657 | -103.0355
9.7695 | -103.0964
9.7732 | -103.1579
9.7769 | -103.2174
9.7806 | -103.2674
9.7844 | -103.3147
9.7881 | -103.3526
9.7918 | -103.3823
9.7955 | -103.4015
9.7993 | -103.4087
9.803 | -103.4091
9.8067 | -103.4044
9.8104 | -103.3889
9.8142 | -103.3561
9.8179 | -103.3136
9.8216 | -103.2648
9.8253 | -103.2116
9.8291 | -103.1546
9.8328 | -103.0997
9.8365 | -103.0468
9.8402 | -102.995
9.844 | -102.9439
9.8477 | -102.8946
9.8514 | -102.8508
9.8552 | -102.8068
9.8589 | -102.7694
9.8626 | -102.7361
9.8663 | -102.7049
9.8701 | -102.6797
9.8738 | -102.6557
9.8775 | -102.6303
9.8812 | -102.6021
9.885 | -102.573
9.8887 | -102.5371
9.8924 | -102.4928
9.8961 | -102.4447
9.8999 | -102.3903
9.9036 | -102.3289
9.9073 | -102.2573
9.911 | -102.1769
9.9148 | -102.0884
9.9185 | -101.9944
9.9222 | -101.8976
9.9259 | -101.7941
9.9297 | -101.6836
9.9334 | -101.5678
9.9371 | -101.4462
9.9408 | -101.3195
9.9446 | -101.1941
9.9483 | -101.074
9.952 | -100.9551
9.9558 | -100.8425
9.9595 | -100.7354
9.9632 | -100.635
9.9669 | -100.5398
9.9707 | -100.4537
9.9744 | -100.3748
9.9781 | -100.3024
9.9818 | -100.2378
9.9856 | -100.182
9.9893 | -100.1325
9.993 | -100.0898
9.9967 | -100.056
10.0005 | -100.0321
10.0042 | -100.0156
10.0079 | -100.0053
10.0116 | -100.0063
10.0154 | -100.0168
10.0191 | -100.0374
10.0228 | -100.0728
10.0265 | -100.1212
10.0303 | -100.1794
10.034 | -100.2513
10.0377 | -100.3383
10.0415 | -100.4362
10.0452 | -100.5455
10.0489 | -100.6638
10.0526 | -100.7889
10.0564 | -100.9253
10.0601 | -101.0726
10.0638 | -101.2238
10.0675 | -101.3754
10.0713 | -101.5315
10.075 | -101.6884
10.0787 | -101.8441
10.0824 | -101.9951
10.0862 | -102.1395
10.0899 | -102.2673
10.0936 | -102.3737
10.0973 | -102.4686
10.1011 | -102.5474
10.1048 | -102.6164
10.1085 | -102.674
10.1122 | -102.7205
10.116 | -102.7536
10.1197 | -102.7784
10.1234 | -102.7992
10.1271 | -102.8172
10.1309 | -102.8296
10.1346 | -102.8396
10.1383 | -102.8515
10.1421 | -102.8597
10.1458 | -102.8689
10.1495 | -102.8877
10.1532 | -102.91
10.157 | -102.9301
10.1607 | -102.95
10.1644 | -102.971
10.1681 | -102.9905
10.1719 | -103.0102
10.1756 | -103.033
10.1793 | -103.0534
10.183 | -103.073
10.1868 | -103.0935
10.1905 | -103.1155
10.1942 | -103.1386
10.1979 | -103.1605
10.2017 | -103.1779
10.2054 | -103.1871
10.2091 | -103.1934
10.2128 | -103.1929
10.2166 | -103.1996
10.2203 | -103.2038
10.224 | -103.2068
10.2277 | -103.2113
10.2315 | -103.2156
10.2352 | -103.2185
10.2389 | -103.2203
10.2427 | -103.2214
10.2464 | -103.2242
10.2501 | -103.2268
10.2538 | -103.2226
10.2576 | -103.215
10.2613 | -103.2103
10.265 | -103.2085
10.2687 | -103.2044
10.2725 | -103.1967
10.2762 | -103.1913
10.2799 | -103.1893
10.2836 | -103.1884
10.2874 | -103.1874
10.2911 | -103.1873
10.2948 | -103.1858
10.2985 | -103.1784
10.3023 | -103.1737
10.306 | -103.1695
10.3097 | -103.1641
10.3134 | -103.1588
10.3172 | -103.1534
10.3209 | -103.1475
10.3246 | -103.1418
10.3284 | -103.1353
10.3321 | -103.128
10.3358 | -103.1197
10.3395 | -103.1128
10.3433 | -103.1091
10.347 | -103.1127
10.3507 | -103.1185
10.3544 | -103.1283
10.3582 | -103.1436
10.3619 | -103.1634
10.3656 | -103.1801
10.3693 | -103.1954
10.3731 | -103.2151
10.3768 | -103.2373
10.3805 | -103.2636
10.3842 | -103.2905
10.388 | -103.316
10.3917 | -103.3433
10.3954 | -103.3729
10.3991 | -103.3975
10.4029 | -103.4202
10.4066 | -103.4434
10.4103 | -103.4618
10.414 | -103.472
10.4178 | -103.4764
10.4215 | -103.4826
10.4252 | -103.4876
10.429 | -103.4838
10.4327 | -103.4757
10.4364 | -103.4624
10.4401 | -103.4465
10.4439 | -103.4292
10.4476 | -103.4069
10.4513 | -103.3809
10.455 | -103.3516
10.4588 | -103.3223
10.4625 | -103.2945
10.4662 | -103.2677
10.4699 | -103.2395
10.4737 | -103.2121
10.4774 | -103.1866
10.4811 | -103.1621
10.4848 | -103.1379
10.4886 | -103.1154
10.4923 | -103.0975
10.496 | -103.0805
10.4997 | -103.0653
10.5035 | -103.0557
10.5072 | -103.0491
10.5109 | -103.045
10.5146 | -103.0451
10.5184 | -103.0542
10.5221 | -103.0674
10.5258 | -103.0767
10.5296 | -103.0869
10.5333 | -103.094
10.537 | -103.1057
10.5407 | -103.122
10.5445 | -103.1392
10.5482 | -103.153
10.5519 | -103.1666
10.5556 | -103.1803
10.5594 | -103.1933
10.5631 | -103.2053
10.5668 | -103.2153
10.5705 | -103.2212
10.5743 | -103.22
10.578 | -103.2131
10.5817 | -103.2078
10.5854 | -103.2017
10.5892 | -103.2002
10.5929 | -103.197
10.5966 | -103.1933
10.6003 | -103.191
10.6041 | -103.1906
10.6078 | -103.1939
10.6115 | -103.195
10.6152 | -103.1926
10.619 | -103.1915
10.6227 | -103.1928
10.6264 | -103.1927
10.6302 | -103.1928
10.6339 | -103.1975
10.6376 | -103.2042
10.6413 | -103.2085
10.6451 | -103.2083
10.6488 | -103.2084
10.6525 | -103.2122
10.6562 | -103.2144
10.66 | -103.2184
10.6637 | -103.2213
10.6674 | -103.2215
10.6711 | -103.2204
10.6749 | -103.2189
10.6786 | -103.2174
10.6823 | -103.2136
10.686 | -103.2103
10.6898 | -103.2041
10.6935 | -103.1956
10.6972 | -103.186
10.7009 | -103.1763
10.7047 | -103.1691
10.7084 | -103.1607
10.7121 | -103.1573
10.7158 | -103.1626
10.7196 | -103.1704
10.7233 | -103.1832
10.727 | -103.2044
10.7308 | -103.2283
10.7345 | -103.256
10.7382 | -103.287
10.7419 | -103.317
10.7457 | -103.3524
10.7494 | -103.3962
10.7531 | -103.4324
10.7568 | -103.4707
10.7606 | -103.5059
10.7643 | -103.5368
10.768 | -103.567
10.7717 | -103.5939
10.7755 | -103.6131
10.7792 | -103.6285
10.7829 | -103.64
10.7866 | -103.6417
10.7904 | -103.6391
10.7941 | -103.6351
10.7978 | -103.6303
10.8015 | -103.6242
10.8053 | -103.6132
10.809 | -103.5944
10.8127 | -103.579
10.8164 | -103.5637
10.8202 | -103.5486
10.8239 | -103.533
10.8276 | -103.513
10.8314 | -103.4937
10.8351 | -103.4778
10.8388 | -103.4639
10.8425 | -103.4499
10.8463 | -103.4348
10.85 | -103.4208
10.8537 | -103.4107
10.8574 | -103.4046
10.8612 | -103.3981
10.8649 | -103.3962
10.8686 | -103.3927
10.8723 | -103.3873
10.8761 | -103.3851
10.8798 | -103.3883
10.8835 | -103.3942
10.8872 | -103.4079
10.891 | -103.4216
10.8947 | -103.4362
10.8984 | -103.4556
10.9021 | -103.4759
10.9059 | -103.4936
10.9096 | -103.5137
10.9133 | -103.5382
10.9171 | -103.558
10.9208 | -103.5789
10.9245 | -103.598
10.9282 | -103.6126
10.932 | -103.6212
10.9357 | -103.6276
10.9394 | -103.6288
10.9431 | -103.6229
10.9469 | -103.6119
10.9506 | -103.5888
10.9543 | -103.5537
10.958 | -103.5143
10.9618 | -103.4711
10.9655 | -103.4264
10.9692 | -103.3789
10.9729 | -103.3261
10.9767 | -103.2684
10.9804 | -103.2085
10.9841 | -103.1513
10.9878 | -103.1004
10.9916 | -103.0498
10.9953 | -103.0004
10.999 | -102.9548
11.0027 | -102.9129
11.0065 | -102.879
11.0102 | -102.8517
11.0139 | -102.8267
11.0177 | -102.8074
11.0214 | -102.7918
11.0251 | -102.7816
11.0288 | -102.7785
11.0326 | -102.7795
11.0363 | -102.7813
11.04 | -102.7853
11.0437 | -102.7895
11.0475 | -102.7932
11.0512 | -102.8012
11.0549 | -102.8112
11.0586 | -102.823
11.0624 | -102.8384
11.0661 | -102.8572
11.0698 | -102.878
11.0735 | -102.8995
11.0773 | -102.9259
11.081 | -102.9555
11.0847 | -102.9873
11.0884 | -103.022
11.0922 | -103.0562
11.0959 | -103.0868
11.0996 | -103.1201
11.1033 | -103.1557
11.1071 | -103.1916
11.1108 | -103.2287
11.1145 | -103.2653
11.1183 | -103.2966
11.122 | -103.325
11.1257 | -103.3503
11.1294 | -103.3734
11.1332 | -103.3919
11.1369 | -103.4108
11.1406 | -103.4256
11.1443 | -103.4364
11.1481 | -103.4446
11.1518 | -103.4526
11.1555 | -103.461
11.1592 | -103.4686
11.163 | -103.4761
11.1667 | -103.4813
11.1704 | -103.4872
11.1741 | -103.4993
11.1779 | -103.5103
11.1816 | -103.5203
11.1853 | -103.5292
11.189 | -103.5402
11.1928 | -103.5527
11.1965 | -103.5706
11.2002 | -103.5869
11.204 | -103.6035
11.2077 | -103.6184
11.2114 | -103.6307
11.2151 | -103.6445
11.2189 | -103.6546
11.2226 | -103.6603
11.2263 | -103.6663
11.23 | -103.6682
11.2338 | -103.6691
11.2375 | -103.6688
11.2412 | -103.6647
11.2449 | -103.6586
11.2487 | -103.6516
11.2524 | -103.6477
11.2561 | -103.6434
11.2598 | -103.6346
11.2636 | -103.6264
11.2673 | -103.6185
11.271 | -103.6114
11.2747 | -103.6049
11.2785 | -103.6035
11.2822 | -103.6029
11.2859 | -103.6141
11.2896 | -103.6164
11.2934 | -103.6231
11.2971 | -103.6322
11.3008 | -103.6405
11.3046 | -103.6535
11.3083 | -103.6669
11.312 | -103.6786
11.3157 | -103.693
11.3195 | -103.7073
11.3232 | -103.7226
11.3269 | -103.7419
11.3306 | -103.7614
11.3344 | -103.7777
11.3381 | -103.7891
11.3418 | -103.8029
11.3455 | -103.8174
11.3493 | -103.8278
11.353 | -103.833
11.3567 | -103.8346
11.3604 | -103.8336
11.3642 | -103.8289
11.3679 | -103.8233
11.3716 | -103.821
11.3753 | -103.8131
11.3791 | -103.799
11.3828 | -103.7877
11.3865 | -103.7799
11.3902 | -103.7691
11.394 | -103.7563
11.3977 | -103.7418
11.4014 | -103.7266
11.4052 | -103.7105
11.4089 | -103.6918
11.4126 | -103.6723
11.4163 | -103.6523
11.4201 | -103.6331
11.4238 | -103.6099
11.4275 | -103.5855
11.4312 | -103.561
11.435 | -103.5342
11.4387 | -103.5095
11.4424 | -103.4844
11.4461 | -103.4599
11.4499 | -103.4387
11.4536 | -103.42
11.4573 | -103.4043
11.461 | -103.3877
11.4648 | -103.3758
11.4685 | -103.367
11.4722 | -103.3589
11.4759 | -103.3534
11.4797 | -103.3511
11.4834 | -103.3516
11.4871 | -103.3516
11.4908 | -103.3528
11.4946 | -103.3566
11.4983 | -103.3656
11.502 | -103.3807
11.5058 | -103.3989
11.5095 | -103.418
11.5132 | -103.439
11.5169 | -103.4603
11.5207 | -103.4832
11.5244 | -103.5089
11.5281 | -103.5313
11.5318 | -103.5534
11.5356 | -103.5737
11.5393 | -103.591
11.543 | -103.6084
11.5467 | -103.6208
11.5505 | -103.6278
11.5542 | -103.6333
11.5579 | -103.6417
11.5616 | -103.6502
11.5654 | -103.6558
11.5691 | -103.6595
11.5728 | -103.6658
11.5765 | -103.672
11.5803 | -103.6786
11.584 | -103.6864
11.5877 | -103.6947
11.5914 | -103.7067
11.5952 | -103.7186
11.5989 | -103.7294
11.6026 | -103.7451
11.6064 | -103.7624
11.6101 | -103.7799
11.6138 | -103.7991
11.6175 | -103.8192
11.6213 | -103.8415
11.625 | -103.8644
11.6287 | -103.8887
11.6324 | -103.9108
11.6362 | -103.9338
11.6399 | -103.9565
11.6436 | -103.9778
11.6473 | -103.9991
11.6511 | -104.0195
11.6548 | -104.0388
11.6585 | -104.0527
11.6622 | -104.0623
11.666 | -104.069
11.6697 | -104.0718
11.6734 | -104.0747
11.6771 | -104.0772
11.6809 | -104.0779
11.6846 | -104.0763
11.6883 | -104.0748
11.6921 | -104.0739
11.6958 | -104.0751
11.6995 | -104.0781
11.7032 | -104.0746
11.707 | -104.0682
11.7107 | -104.067
11.7144 | -104.0696
11.7181 | -104.0698
11.7219 | -104.0645
11.7256 | -104.0582
11.7293 | -104.0547
11.733 | -104.0543
11.7368 | -104.0544
11.7405 | -104.0507
11.7442 | -104.0478
11.7479 | -104.0432
11.7517 | -104.0383
11.7554 | -104.0335
11.7591 | -104.0296
11.7628 | -104.026
11.7666 | -104.0228
11.7703 | -104.0164
11.774 | -104.0103
11.7777 | -104.0045
11.7815 | -103.9982
11.7852 | -103.993
11.7889 | -103.9855
11.7927 | -103.9757
11.7964 | -103.9657
11.8001 | -103.9564
11.8038 | -103.9436
11.8076 | -103.9284
11.8113 | -103.9137
11.815 | -103.8959
11.8187 | -103.8749
11.8225 | -103.8412
11.8262 | -103.8174
11.8299 | -103.7944
11.8336 | -103.7701
11.8374 | -103.7455
11.8411 | -103.7214
11.8448 | -103.6979
11.8485 | -103.6767
11.8523 | -103.6583
11.856 | -103.6451
11.8597 | -103.6336
11.8634 | -103.6241
11.8672 | -103.6189
11.8709 | -103.6189
11.8746 | -103.6294
11.8783 | -103.6428
11.8821 | -103.6556
11.8858 | -103.6705
11.8895 | -103.6896
11.8933 | -103.7132
11.897 | -103.7375
11.9007 | -103.7619
11.9044 | -103.7856
11.9082 | -103.8126
11.9119 | -103.8383
11.9156 | -103.8628
11.9193 | -103.8863
11.9231 | -103.9068
11.9268 | -103.9248
11.9305 | -103.9421
11.9342 | -103.9597
11.938 | -103.9711
11.9417 | -103.975
11.9454 | -103.9793
11.9491 | -103.9844
11.9529 | -103.9897
11.9566 | -103.9911
11.9603 | -103.9905
11.964 | -103.9909
11.9678 | -103.9916
11.9715 | -103.9853
11.9752 | -103.9794
11.9789 | -103.9747
11.9827 | -103.9647
11.9864 | -103.9536
11.9901 | -103.9432
11.9939 | -103.9362
11.9976 | -103.9269
12.0013 | -103.9154
12.005 | -103.9036
12.0088 | -103.893
12.0125 | -103.8814
12.0162 | -103.8713
12.0199 | -103.8637
12.0237 | -103.8582
12.0274 | -103.8581
12.0311 | -103.8614
12.0348 | -103.87
12.0386 | -103.8796
12.0423 | -103.8904
12.046 | -103.9055
12.0497 | -103.9239
12.0535 | -103.943
12.0572 | -103.9647
12.0609 | -103.9879
12.0646 | -104.0126
12.0684 | -104.0369
12.0721 | -104.0627
12.0758 | -104.0888
12.0795 | -104.1141
12.0833 | -104.1353
12.087 | -104.1562
12.0907 | -104.1754
12.0945 | -104.1924
12.0982 | -104.2061
12.1019 | -104.2193
12.1056 | -104.2269
12.1094 | -104.228
12.1131 | -104.228
12.1168 | -104.2295
12.1205 | -104.2235
12.1243 | -104.2145
12.128 | -104.2055
12.1317 | -104.1945
12.1354 | -104.1844
12.1392 | -104.1723
12.1429 | -104.1601
12.1466 | -104.1554
12.1503 | -104.1465
12.1541 | -104.1311
12.1578 | -104.1195
12.1615 | -104.1134
12.1652 | -104.103
12.169 | -104.0883
12.1727 | -104.0723
12.1764 | -104.0564
12.1801 | -104.0423
12.1839 | -104.0253
12.1876 | -104.0065
12.1913 | -103.9851
12.1951 | -103.963
12.1988 | -103.9368
12.2025 | -103.9095
12.2062 | -103.8834
12.21 | -103.8553
12.2137 | -103.8293
12.2174 | -103.806
12.2211 | -103.7804
12.2249 | -103.756
12.2286 | -103.7369
12.2323 | -103.7172
12.236 | -103.6973
12.2398 | -103.6791
12.2435 | -103.6598
12.2472 | -103.6411
12.2509 | -103.6254
12.2547 | -103.6153
12.2584 | -103.6066
12.2621 | -103.5999
12.2658 | -103.5933
12.2696 | -103.5923
12.2733 | -103.5922
12.277 | -103.5939
12.2808 | -103.6008
12.2845 | -103.6097
12.2882 | -103.6236
12.2919 | -103.6384
12.2957 | -103.6524
12.2994 | -103.666
12.3031 | -103.6791
12.3068 | -103.6902
12.3106 | -103.7007
12.3143 | -103.7092
12.318 | -103.7154
12.3217 | -103.7151
12.3255 | -103.7167
12.3292 | -103.719
12.3329 | -103.7139
12.3366 | -103.702
12.3404 | -103.6923
12.3441 | -103.6821
12.3478 | -103.6674
12.3515 | -103.6513
12.3553 | -103.6305
12.359 | -103.6093
12.3627 | -103.5892
12.3664 | -103.5706
12.3702 | -103.5522
12.3739 | -103.5308
12.3776 | -103.5147
12.3814 | -103.5004
12.3851 | -103.4853
12.3888 | -103.4716
12.3925 | -103.4615
12.3963 | -103.4546
12.4 | -103.4502
12.4037 | -103.45
12.4074 | -103.4508
12.4112 | -103.4546
12.4149 | -103.4634
12.4186 | -103.4734
12.4223 | -103.4883
12.4261 | -103.5045
12.4298 | -103.5202
12.4335 | -103.5363
12.4372 | -103.5531
12.441 | -103.571
12.4447 | -103.5909
12.4484 | -103.6061
12.4521 | -103.6226
12.4559 | -103.6369
12.4596 | -103.6505
12.4633 | -103.6636
12.467 | -103.6713
12.4708 | -103.6766
12.4745 | -103.6826
12.4782 | -103.6867
12.482 | -103.6887
12.4857 | -103.6908
12.4894 | -103.6894
12.4931 | -103.6838
12.4969 | -103.6747
12.5006 | -103.6672
12.5043 | -103.6547
12.508 | -103.6415
12.5118 | -103.6287
12.5155 | -103.6155
12.5192 | -103.6026
12.5229 | -103.5901
12.5267 | -103.5804
12.5304 | -103.5698
12.5341 | -103.5619
12.5378 | -103.5563
12.5416 | -103.5494
12.5453 | -103.5473
12.549 | -103.5479
12.5527 | -103.5484
12.5565 | -103.5497
12.5602 | -103.5509
12.5639 | -103.5547
12.5676 | -103.5602
12.5714 | -103.5638
12.5751 | -103.5688
12.5788 | -103.5726
12.5826 | -103.5762
12.5863 | -103.5786
12.59 | -103.5797
12.5937 | -103.5784
12.5975 | -103.5792
12.6012 | -103.5819
12.6049 | -103.5823
12.6086 | -103.5835
12.6124 | -103.5839
12.6161 | -103.5815
12.6198 | -103.5785
12.6235 | -103.574
12.6273 | -103.5666
12.631 | -103.5589
12.6347 | -103.5506
12.6384 | -103.5413
12.6422 | -103.5311
12.6459 | -103.5225
12.6496 | -103.5167
12.6533 | -103.512
12.6571 | -103.5069
12.6608 | -103.501
12.6645 | -103.4995
12.6683 | -103.5022
12.672 | -103.5073
12.6757 | -103.516
12.6794 | -103.5272
12.6832 | -103.5369
12.6869 | -103.5499
12.6906 | -103.5657
12.6943 | -103.5858
12.6981 | -103.6055
12.7018 | -103.6234
12.7055 | -103.6425
12.7092 | -103.6647
12.713 | -103.6834
12.7167 | -103.7004
12.7204 | -103.7133
12.7241 | -103.7245
12.7279 | -103.7332
12.7316 | -103.7379
12.7353 | -103.7394
12.739 | -103.737
12.7428 | -103.7315
12.7465 | -103.7208
12.7502 | -103.708
12.7539 | -103.6951
12.7577 | -103.6793
12.7614 | -103.6613
12.7651 | -103.643
12.7689 | -103.6231
12.7726 | -103.6051
12.7763 | -103.5862
12.78 | -103.566
12.7838 | -103.5448
12.7875 | -103.5288
12.7912 | -103.5152
12.7949 | -103.5041
12.7987 | -103.4947
12.8024 | -103.4875
12.8061 | -103.4821
12.8098 | -103.4815
12.8136 | -103.4842
12.8173 | -103.488
12.821 | -103.4926
12.8247 | -103.4971
12.8285 | -103.4981
12.8322 | -103.502
12.8359 | -103.5101
12.8396 | -103.5172
12.8434 | -103.5295
12.8471 | -103.5413
12.8508 | -103.5521
12.8545 | -103.5637
12.8583 | -103.5758
12.862 | -103.5874
12.8657 | -103.5976
12.8695 | -103.6073
12.8732 | -103.6142
12.8769 | -103.6192
12.8806 | -103.6267
12.8844 | -103.6352
12.8881 | -103.6437
12.8918 | -103.6501
12.8955 | -103.6616
12.8993 | -103.6776
12.903 | -103.6881
12.9067 | -103.703
12.9104 | -103.7174
12.9142 | -103.7311
12.9179 | -103.7476
12.9216 | -103.7669
12.9253 | -103.7881
12.9291 | -103.8052
12.9328 | -103.8218
12.9365 | -103.8413
12.9402 | -103.8608
12.944 | -103.8746
12.9477 | -103.889
12.9514 | -103.9056
12.9551 | -103.9221
12.9589 | -103.9357
12.9626 | -103.9462
12.9663 | -103.9521
12.9701 | -103.9562
12.9738 | -103.9583
12.9775 | -103.9588
12.9812 | -103.9596
12.985 | -103.9567
12.9887 | -103.9498
12.9924 | -103.9445
12.9961 | -103.9336
12.9999 | -103.9207
13.0036 | -103.9031
13.0073 | -103.8852
13.011 | -103.8696
13.0148 | -103.8541
13.0185 | -103.8353
13.0222 | -103.8143
13.0259 | -103.7939
13.0297 | -103.7688
13.0334 | -103.7475
13.0371 | -103.7273
13.0408 | -103.7088
13.0446 | -103.6902
13.0483 | -103.6741
13.052 | -103.6582
13.0557 | -103.641
13.0595 | -103.626
13.0632 | -103.6113
13.0669 | -103.5946
13.0707 | -103.5802
13.0744 | -103.5661
13.0781 | -103.5552
13.0818 | -103.5431
13.0856 | -103.5306
13.0893 | -103.5199
13.093 | -103.5113
13.0967 | -103.5007
13.1005 | -103.4904
13.1042 | -103.4792
13.1079 | -103.4672
13.1116 | -103.4563
13.1154 | -103.4446
13.1191 | -103.4335
13.1228 | -103.4232
13.1265 | -103.409
13.1303 | -103.395
13.134 | -103.3821
13.1377 | -103.3692
13.1414 | -103.3533
13.1452 | -103.3352
13.1489 | -103.317
13.1526 | -103.3042
13.1563 | -103.2898
13.1601 | -103.2678
13.1638 | -103.245
13.1675 | -103.2258
13.1713 | -103.2066
13.175 | -103.1922
13.1787 | -103.1823
13.1824 | -103.1703
13.1862 | -103.1592
13.1899 | -103.1471
13.1936 | -103.1387
13.1973 | -103.1304
13.2011 | -103.1199
13.2048 | -103.1101
13.2085 | -103.1015
13.2122 | -103.0924
13.216 | -103.0839
13.2197 | -103.0749
13.2234 | -103.0654
13.2271 | -103.058
13.2309 | -103.0536
13.2346 | -103.0514
13.2383 | -103.0487
13.242 | -103.0511
13.2458 | -103.0585
13.2495 | -103.0676
13.2532 | -103.0782
13.257 | -103.0926
13.2607 | -103.1071
13.2644 | -103.1225
13.2681 | -103.1368
13.2719 | -103.1538
13.2756 | -103.1734
13.2793 | -103.1924
13.283 | -103.2074
13.2868 | -103.2205
13.2905 | -103.2365
13.2942 | -103.2515
13.2979 | -103.2642
13.3017 | -103.2744
13.3054 | -103.2846
13.3091 | -103.2946
13.3128 | -103.2992
13.3166 | -103.2998
13.3203 | -103.2988
13.324 | -103.2934
13.3277 | -103.2851
13.3315 | -103.2754
13.3352 | -103.267
13.3389 | -103.2595
13.3426 | -103.2457
13.3464 | -103.2324
13.3501 | -103.2247
13.3538 | -103.2151
13.3576 | -103.2056
13.3613 | -103.1967
13.365 | -103.1871
13.3687 | -103.1807
13.3725 | -103.1776
13.3762 | -103.1792
13.3799 | -103.1851
13.3836 | -103.1951
13.3874 | -103.2085
13.3911 | -103.2253
13.3948 | -103.2478
13.3985 | -103.2731
13.4023 | -103.3
13.406 | -103.3309
13.4097 | -103.3645
13.4134 | -103.3981
13.4172 | -103.4341
13.4209 | -103.4685
13.4246 | -103.5064
13.4283 | -103.549
13.4321 | -103.5917
13.4358 | -103.6325
13.4395 | -103.6726
13.4432 | -103.7124
13.447 | -103.7525
13.4507 | -103.792
13.4544 | -103.8327
13.4582 | -103.8729
13.4619 | -103.9141
13.4656 | -103.9513
13.4693 | -103.9885
13.4731 | -104.0246
13.4768 | -104.0605
13.4805 | -104.0938
13.4842 | -104.122
13.488 | -104.1501
13.4917 | -104.1752
13.4954 | -104.1954
13.4991 | -104.2129
13.5029 | -104.2279
13.5066 | -104.245
13.5103 | -104.263
13.514 | -104.2774
13.5178 | -104.2895
13.5215 | -104.2979
13.5252 | -104.3037
13.5289 | -104.3045
13.5327 | -104.3004
13.5364 | -104.301
13.5401 | -104.3003
13.5439 | -104.2959
13.5476 | -104.2868
13.5513 | -104.2736
13.555 | -104.2636
13.5588 | -104.2554
13.5625 | -104.2426
13.5662 | -104.2279
13.5699 | -104.2107
13.5737 | -104.1964
13.5774 | -104.1801
13.5811 | -104.1686
13.5848 | -104.1573
13.5886 | -104.1421
13.5923 | -104.1289
13.596 | -104.1213
13.5997 | -104.1137
13.6035 | -104.1072
13.6072 | -104.0981
13.6109 | -104.0883
13.6146 | -104.0817
13.6184 | -104.073
13.6221 | -104.0641
13.6258 | -104.0555
13.6295 | -104.0498
13.6333 | -104.0463
13.637 | -104.0392
13.6407 | -104.0313
13.6445 | -104.0267
13.6482 | -104.0239
13.6519 | -104.0208
13.6556 | -104.0177
13.6594 | -104.012
13.6631 | -104.0091
13.6668 | -104.0062
13.6705 | -104.0024
13.6743 | -104.0015
13.678 | -104.0032
13.6817 | -104.0046
13.6854 | -104.0063
13.6892 | -104.0116
13.6929 | -104.0202
13.6966 | -104.0317
13.7003 | -104.046
13.7041 | -104.0612
13.7078 | -104.0783
13.7115 | -104.0985
13.7152 | -104.1204
13.719 | -104.141
13.7227 | -104.1649
13.7264 | -104.1938
13.7301 | -104.2231
13.7339 | -104.2504
13.7376 | -104.2773
13.7413 | -104.309
13.7451 | -104.3447
13.7488 | -104.3816
13.7525 | -104.4156
13.7562 | -104.4483
13.76 | -104.4821
13.7637 | -104.5159
13.7674 | -104.5512
13.7711 | -104.5808
13.7749 | -104.6075
13.7786 | -104.6259
13.7823 | -104.6433
13.786 | -104.6633
13.7898 | -104.6763
13.7935 | -104.684
13.7972 | -104.6924
13.8009 | -104.6989
13.8047 | -104.7032
13.8084 | -104.703
13.8121 | -104.7016
13.8158 | -104.7002
13.8196 | -104.695
13.8233 | -104.6903
13.827 | -104.6809
13.8307 | -104.6712
13.8345 | -104.6677
13.8382 | -104.6661
13.8419 | -104.6612
13.8457 | -104.6548
13.8494 | -104.6488
13.8531 | -104.6481
13.8568 | -104.6528
13.8606 | -104.6588
13.8643 | -104.6633
13.868 | -104.6684
13.8717 | -104.6742
13.8755 | -104.681
13.8792 | -104.6892
13.8829 | -104.6969
13.8866 | -104.7028
13.8904 | -104.7083
13.8941 | -104.7123
13.8978 | -104.7122
13.9015 | -104.7139
13.9053 | -104.7156
13.909 | -104.7152
13.9127 | -104.7169
13.9164 | -104.7186
13.9202 | -104.716
13.9239 | -104.7106
13.9276 | -104.7042
13.9313 | -104.6952
13.9351 | -104.6873
13.9388 | -104.6769
13.9425 | -104.6657
13.9463 | -104.6522
13.95 | -104.6408
13.9537 | -104.63
13.9574 | -104.6188
13.9612 | -104.598
13.9649 | -104.561
13.9686 | -104.526
13.9723 | -104.4858
13.9761 | -104.4581
13.9798 | -104.4288
13.9835 | -104.3974
13.9872 | -104.3618
13.991 | -104.3262
13.9947 | -104.2844
13.9984 | -104.2379
14.0021 | -104.1882
14.0059 | -104.1372
14.0096 | -104.086
14.0133 | -104.029
14.017 | -103.9683
14.0208 | -103.9081
14.0245 | -103.8479
14.0282 | -103.7928
14.032 | -103.7358
14.0357 | -103.6784
14.0394 | -103.6265
14.0431 | -103.5799
14.0469 | -103.5385
14.0506 | -103.5019
14.0543 | -103.4687
14.058 | -103.4413
14.0618 | -103.4194
14.0655 | -103.4059
14.0692 | -103.4016
14.0729 | -103.4044
14.0767 | -103.4141
14.0804 | -103.4277
14.0841 | -103.4484
14.0878 | -103.4762
14.0916 | -103.5107
14.0953 | -103.5497
14.099 | -103.5943
14.1027 | -103.6393
14.1065 | -103.6862
14.1102 | -103.7354
14.1139 | -103.7863
14.1176 | -103.8387
14.1214 | -103.8929
14.1251 | -103.9479
14.1288 | -104.0039
14.1326 | -104.0588
14.1363 | -104.1115
14.14 | -104.1591
14.1437 | -104.204
14.1475 | -104.2494
14.1512 | -104.2925
14.1549 | -104.3358
14.1586 | -104.3769
14.1624 | -104.4209
14.1661 | -104.4704
14.1698 | -104.5109
14.1735 | -104.539
14.1773 | -104.5623
14.181 | -104.5852
14.1847 | -104.607
14.1884 | -104.6257
14.1922 | -104.6429
14.1959 | -104.6592
14.1996 | -104.6797
14.2033 | -104.6996
14.2071 | -104.7205
14.2108 | -104.7404
14.2145 | -104.7607
14.2182 | -104.7837
14.222 | -104.8066
14.2257 | -104.8291
14.2294 | -104.8529
14.2332 | -104.8739
14.2369 | -104.8976
14.2406 | -104.9216
14.2443 | -104.9444
14.2481 | -104.9664
14.2518 | -104.9874
14.2555 | -105.0109
14.2592 | -105.0313
14.263 | -105.0494
14.2667 | -105.0661
14.2704 | -105.082
14.2741 | -105.0952
14.2779 | -105.103
14.2816 | -105.111
14.2853 | -105.1226
14.289 | -105.1312
14.2928 | -105.1353
14.2965 | -105.1373
14.3002 | -105.1375
14.3039 | -105.1388
14.3077 | -105.1422
14.3114 | -105.1424
14.3151 | -105.1363
14.3188 | -105.1271
14.3226 | -105.118
14.3263 | -105.1098
14.33 | -105.1
14.3338 | -105.089
14.3375 | -105.0766
14.3412 | -105.0657
14.3449 | -105.0582
14.3487 | -105.0518
14.3524 | -105.0439
14.3561 | -105.0363
14.3598 | -105.0291
14.3636 | -105.0229
14.3673 | -105.0193
14.371 | -105.0204
14.3747 | -105.0185
14.3785 | -105.016
14.3822 | -105.0179
14.3859 | -105.0204
14.3896 | -105.0248
14.3934 | -105.0332
14.3971 | -105.044
14.4008 | -105.0547
14.4045 | -105.0648
14.4083 | -105.077
14.412 | -105.0902
14.4157 | -105.1049
14.4194 | -105.1172
14.4232 | -105.1279
14.4269 | -105.1415
14.4306 | -105.1531
14.4344 | -105.1649
14.4381 | -105.1789
14.4418 | -105.1921
14.4455 | -105.2044
14.4493 | -105.2152
14.453 | -105.2253
14.4567 | -105.2348
14.4604 | -105.2417
14.4642 | -105.2482
14.4679 | -105.2557
14.4716 | -105.2625
14.4753 | -105.2656
14.4791 | -105.2667
14.4828 | -105.2679
14.4865 | -105.2614
14.4902 | -105.2514
14.494 | -105.2396
14.4977 | -105.2265
14.5014 | -105.2126
14.5051 | -105.1973
14.5089 | -105.1771
14.5126 | -105.1508
14.5163 | -105.127
14.52 | -105.107
14.5238 | -105.088
14.5275 | -105.07
14.5312 | -105.0491
14.535 | -105.0241
14.5387 | -105.0006
14.5424 | -104.9811
14.5461 | -104.9635
14.5499 | -104.9443
14.5536 | -104.9257
14.5573 | -104.9065
14.561 | -104.8899
14.5648 | -104.877
14.5685 | -104.8681
14.5722 | -104.8617
14.5759 | -104.8578
14.5797 | -104.8518
14.5834 | -104.8485
14.5871 | -104.8474
14.5908 | -104.8477
14.5946 | -104.8515
14.5983 | -104.8552
14.602 | -104.858
14.6057 | -104.8616
14.6095 | -104.8657
14.6132 | -104.8719
14.6169 | -104.8779
14.6207 | -104.8835
14.6244 | -104.8873
14.6281 | -104.8903
14.6318 | -104.8988
14.6356 | -104.8998
14.6393 | -104.9017
14.643 | -104.9028
14.6467 | -104.9022
14.6505 | -104.9013
14.6542 | -104.9
14.6579 | -104.8987
14.6616 | -104.8971
14.6654 | -104.8967
14.6691 | -104.8965
14.6728 | -104.8963
14.6765 | -104.8936
14.6803 | -104.8899
14.684 | -104.886
14.6877 | -104.8826
14.6914 | -104.8806
14.6952 | -104.8807
14.6989 | -104.8813
14.7026 | -104.8809
14.7063 | -104.8797
14.7101 | -104.8748
14.7138 | -104.8675
14.7175 | -104.8676
14.7213 | -104.8673
14.725 | -104.8692
14.7287 | -104.8742
14.7324 | -104.877
14.7362 | -104.8826
14.7399 | -104.8881
14.7436 | -104.8965
14.7473 | -104.9091
14.7511 | -104.9213
14.7548 | -104.9321
14.7585 | -104.9449
14.7622 | -104.9602
14.766 | -104.9774
14.7697 | -104.9958
14.7734 | -105.0134
14.7771 | -105.0297
14.7809 | -105.0444
14.7846 | -105.0613
14.7883 | -105.0774
14.792 | -105.0948
14.7958 | -105.1113
14.7995 | -105.1264
14.8032 | -105.1387
14.8069 | -105.1499
14.8107 | -105.1618
14.8144 | -105.1748
14.8181 | -105.1861
14.8219 | -105.1972
14.8256 | -105.2082
14.8293 | -105.2191
14.833 | -105.228
14.8368 | -105.2364
14.8405 | -105.2453
14.8442 | -105.247
14.8479 | -105.2549
14.8517 | -105.2613
14.8554 | -105.2637
14.8591 | -105.2644
14.8628 | -105.2644
14.8666 | -105.2634
14.8703 | -105.2611
14.874 | -105.2594
14.8777 | -105.2536
14.8815 | -105.2458
14.8852 | -105.2363
14.8889 | -105.2295
14.8926 | -105.2225
14.8964 | -105.2163
14.9001 | -105.2116
14.9038 | -105.2058
14.9076 | -105.1998
14.9113 | -105.1946
14.915 | -105.1916
14.9187 | -105.1917
14.9225 | -105.1961
14.9262 | -105.1998
14.9299 | -105.195
14.9336 | -105.1898
14.9374 | -105.1865
14.9411 | -105.1814
14.9448 | -105.1762
14.9485 | -105.1723
14.9523 | -105.1655
14.956 | -105.1594
14.9597 | -105.1505
14.9634 | -105.1438
14.9672 | -105.1389
14.9709 | -105.1328
14.9746 | -105.1248
14.9783 | -105.1129
14.9821 | -105.0984
14.9858 | -105.0842
14.9895 | -105.0684
14.9932 | -105.0505
14.997 | -105.0337
15.0007 | -105.0192
15.0044 | -105.0058
15.0082 | -104.9908
15.0119 | -104.9764
15.0156 | -104.9613
15.0193 | -104.9427
15.0231 | -104.9266
15.0268 | -104.9133
15.0305 | -104.9027
15.0342 | -104.8923
15.038 | -104.879
15.0417 | -104.8631
15.0454 | -104.8526
15.0491 | -104.8519
15.0529 | -104.8475
15.0566 | -104.844
15.0603 | -104.8428
15.064 | -104.8401
15.0678 | -104.8335
15.0715 | -104.8297
15.0752 | -104.8295
15.0789 | -104.8294
15.0827 | -104.8307
15.0864 | -104.8298
15.0901 | -104.8275
15.0938 | -104.8282
15.0976 | -104.8297
15.1013 | -104.8303
15.105 | -104.8289
15.1088 | -104.8277
15.1125 | -104.8265
15.1162 | -104.8261
15.1199 | -104.8258
15.1237 | -104.8271
15.1274 | -104.8273
15.1311 | -104.8291
15.1348 | -104.8339
15.1386 | -104.8372
15.1423 | -104.8429
15.146 | -104.8561
15.1497 | -104.8701
15.1535 | -104.8836
15.1572 | -104.8985
15.1609 | -104.9159
15.1646 | -104.9322
15.1684 | -104.9517
15.1721 | -104.9721
15.1758 | -104.9931
15.1795 | -105.0117
15.1833 | -105.0293
15.187 | -105.0488
15.1907 | -105.0682
15.1944 | -105.09
15.1982 | -105.1118
15.2019 | -105.1336
15.2056 | -105.1522
15.2094 | -105.1663
15.2131 | -105.1781
15.2168 | -105.1886
15.2205 | -105.1962
15.2243 | -105.2006
15.228 | -105.2015
15.2317 | -105.2008
15.2354 | -105.1986
15.2392 | -105.191
15.2429 | -105.1772
15.2466 | -105.1598
15.2503 | -105.1376
15.2541 | -105.1159
15.2578 | -105.0953
15.2615 | -105.069
15.2652 | -105.0405
15.269 | -105.0057
15.2727 | -104.9682
15.2764 | -104.9326
15.2801 | -104.8965
15.2839 | -104.8592
15.2876 | -104.8282
15.2913 | -104.795
15.295 | -104.7576
15.2988 | -104.721
15.3025 | -104.689
15.3062 | -104.6604
15.31 | -104.6342
15.3137 | -104.6116
15.3174 | -104.5912
15.3211 | -104.5738
15.3249 | -104.5568
15.3286 | -104.5439
15.3323 | -104.5333
15.336 | -104.5242
15.3398 | -104.5187
15.3435 | -104.5178
15.3472 | -104.5204
15.3509 | -104.5254
15.3547 | -104.5326
15.3584 | -104.5452
15.3621 | -104.5591
15.3658 | -104.5723
15.3696 | -104.5901
15.3733 | -104.6104
15.377 | -104.6325
15.3807 | -104.6571
15.3845 | -104.6824
15.3882 | -104.7085
15.3919 | -104.7349
15.3956 | -104.7626
15.3994 | -104.7909
15.4031 | -104.8176
15.4068 | -104.8422
15.4106 | -104.8647
15.4143 | -104.8865
15.418 | -104.908
15.4217 | -104.925
15.4255 | -104.941
15.4292 | -104.9548
15.4329 | -104.9614
15.4366 | -104.9627
15.4404 | -104.9577
15.4441 | -104.9542
15.4478 | -104.9507
15.4515 | -104.9415
15.4553 | -104.9298
15.459 | -104.9153
15.4627 | -104.8989
15.4664 | -104.8769
15.4702 | -104.8541
15.4739 | -104.8311
15.4776 | -104.8044
15.4813 | -104.7714
15.4851 | -104.7313
15.4888 | -104.6956
15.4925 | -104.6616
15.4963 | -104.624
15.5 | -104.5852
15.5037 | -104.5492
15.5074 | -104.5134
15.5112 | -104.4768
15.5149 | -104.4439
15.5186 | -104.4131
15.5223 | -104.3774
15.5261 | -104.3431
15.5298 | -104.311
15.5335 | -104.2754
15.5372 | -104.2461
15.541 | -104.2226
15.5447 | -104.2
15.5484 | -104.1763
15.5521 | -104.1553
15.5559 | -104.1344
15.5596 | -104.1138
15.5633 | -104.0938
15.567 | -104.0775
15.5708 | -104.0605
15.5745 | -104.043
15.5782 | -104.0268
15.5819 | -104.012
15.5857 | -103.9977
15.5894 | -103.9818
15.5931 | -103.9649
15.5969 | -103.9506
15.6006 | -103.9371
15.6043 | -103.9235
15.608 | -103.9133
15.6118 | -103.9042
15.6155 | -103.8929
15.6192 | -103.879
15.6229 | -103.8649
15.6267 | -103.8543
15.6304 | -103.8447
15.6341 | -103.8356
15.6378 | -103.8288
15.6416 | -103.8216
15.6453 | -103.814
15.649 | -103.8072
15.6527 | -103.8039
15.6565 | -103.8027
15.6602 | -103.8034
15.6639 | -103.8059
15.6676 | -103.8079
15.6714 | -103.8046
15.6751 | -103.8018
15.6788 | -103.8008
15.6825 | -103.8024
15.6863 | -103.8037
15.69 | -103.806
15.6937 | -103.8088
15.6975 | -103.8059
15.7012 | -103.8055
15.7049 | -103.8092
15.7086 | -103.8161
15.7124 | -103.8177
15.7161 | -103.8154
15.7198 | -103.8126
15.7235 | -103.8124
15.7273 | -103.8106
15.731 | -103.8028
15.7347 | -103.7976
15.7384 | -103.79
15.7422 | -103.7813
15.7459 | -103.7753
15.7496 | -103.7688
15.7533 | -103.7614
15.7571 | -103.7542
15.7608 | -103.7419
15.7645 | -103.7301
15.7682 | -103.718
15.772 | -103.7047
15.7757 | -103.6893
15.7794 | -103.6726
15.7831 | -103.6547
15.7869 | -103.6395
15.7906 | -103.6219
15.7943 | -103.6057
15.7981 | -103.5901
15.8018 | -103.5713
15.8055 | -103.5513
15.8092 | -103.534
15.813 | -103.5165
15.8167 | -103.5
15.8204 | -103.4842
15.8241 | -103.4677
15.8279 | -103.4517
15.8316 | -103.4331
15.8353 | -103.4136
15.839 | -103.3961
15.8428 | -103.3787
15.8465 | -103.3643
15.8502 | -103.3541
15.8539 | -103.3394
15.8577 | -103.3243
15.8614 | -103.31
15.8651 | -103.2932
15.8688 | -103.2783
15.8726 | -103.268
15.8763 | -103.2595
15.88 | -103.2511
15.8837 | -103.2447
15.8875 | -103.2386
15.8912 | -103.2338
15.8949 | -103.2315
15.8987 | -103.2366
15.9024 | -103.245
15.9061 | -103.2515
15.9098 | -103.2524
15.9136 | -103.2598
15.9173 | -103.2739
15.921 | -103.289
15.9247 | -103.3084
15.9285 | -103.3282
15.9322 | -103.3471
15.9359 | -103.3677
15.9396 | -103.3934
15.9434 | -103.4206
15.9471 | -103.4499
15.9508 | -103.4777
15.9545 | -103.5058
15.9583 | -103.5372
15.962 | -103.567
15.9657 | -103.5971
15.9694 | -103.6261
15.9732 | -103.6541
15.9769 | -103.6824
15.9806 | -103.7095
15.9844 | -103.7377
15.9881 | -103.7666
15.9918 | -103.7932
15.9955 | -103.8199
15.9993 | -103.8496
16.003 | -103.879
16.0067 | -103.9063
16.0104 | -103.9333
16.0142 | -103.9549
16.0179 | -103.9763
16.0216 | -103.998
16.0253 | -104.0194
16.0291 | -104.043
16.0328 | -104.0659
16.0365 | -104.0841
16.0402 | -104.1051
16.044 | -104.1239
16.0477 | -104.1417
16.0514 | -104.1582
16.0551 | -104.173
16.0589 | -104.1889
16.0626 | -104.2012
16.0663 | -104.2097
16.07 | -104.2226
16.0738 | -104.2363
16.0775 | -104.2454
16.0812 | -104.2555
16.085 | -104.2655
16.0887 | -104.2745
16.0924 | -104.2815
16.0961 | -104.2872
16.0999 | -104.288
16.1036 | -104.2884
16.1073 | -104.2908
16.111 | -104.2942
16.1148 | -104.2993
16.1185 | -104.3069
16.1222 | -104.3138
16.1259 | -104.3229
16.1297 | -104.3271
16.1334 | -104.3312
16.1371 | -104.3393
16.1408 | -104.3562
16.1446 | -104.3678
16.1483 | -104.3754
16.152 | -104.3846
16.1557 | -104.3928
16.1595 | -104.4001
16.1632 | -104.4124
16.1669 | -104.4254
16.1706 | -104.4376
16.1744 | -104.4495
16.1781 | -104.46
16.1818 | -104.4704
16.1856 | -104.4828
16.1893 | -104.4976
16.193 | -104.512
16.1967 | -104.5272
16.2005 | -104.5427
16.2042 | -104.5546
16.2079 | -104.5656
16.2116 | -104.5765
16.2154 | -104.5842
16.2191 | -104.5914
16.2228 | -104.5984
16.2265 | -104.6067
16.2303 | -104.613
16.234 | -104.619
16.2377 | -104.6269
16.2414 | -104.6334
16.2452 | -104.6389
16.2489 | -104.6411
16.2526 | -104.6411
16.2563 | -104.6418
16.2601 | -104.6434
16.2638 | -104.6456
16.2675 | -104.6482
16.2712 | -104.6504
16.275 | -104.6536
16.2787 | -104.6555
16.2824 | -104.6574
16.2862 | -104.6608
16.2899 | -104.667
16.2936 | -104.6736
16.2973 | -104.6797
16.3011 | -104.6835
16.3048 | -104.6888
16.3085 | -104.6961
16.3122 | -104.7064
16.316 | -104.7173
16.3197 | -104.7289
16.3234 | -104.7433
16.3271 | -104.7607
16.3309 | -104.7805
16.3346 | -104.8006
16.3383 | -104.8215
16.342 | -104.8414
16.3458 | -104.862
16.3495 | -104.8855
16.3532 | -104.9088
16.3569 | -104.9349
16.3607 | -104.9629
16.3644 | -104.9938
16.3681 | -105.0266
16.3718 | -105.0585
16.3756 | -105.088
16.3793 | -105.1202
16.383 | -105.1538
16.3868 | -105.188
16.3905 | -105.225
16.3942 | -105.262
16.3979 | -105.2947
16.4017 | -105.3224
16.4054 | -105.3471
16.4091 | -105.3743
16.4128 | -105.4017
16.4166 | -105.4289
16.4203 | -105.4496
16.424 | -105.4661
16.4277 | -105.4824
16.4315 | -105.4958
16.4352 | -105.5069
16.4389 | -105.5199
16.4426 | -105.5302
16.4464 | -105.5384
16.4501 | -105.5461
16.4538 | -105.5489
16.4575 | -105.5455
16.4613 | -105.5419
16.465 | -105.5379
16.4687 | -105.5326
16.4725 | -105.5274
16.4762 | -105.5206
16.4799 | -105.5125
16.4836 | -105.5074
16.4874 | -105.499
16.4911 | -105.4892
16.4948 | -105.4805
16.4985 | -105.4722
16.5023 | -105.4649
16.506 | -105.4562
16.5097 | -105.4469
16.5134 | -105.4424
16.5172 | -105.439
16.5209 | -105.4339
16.5246 | -105.4296
16.5283 | -105.4309
16.5321 | -105.4333
16.5358 | -105.4346
16.5395 | -105.4325
16.5432 | -105.4304
16.547 | -105.4265
16.5507 | -105.4252
16.5544 | -105.4265
16.5581 | -105.4273
16.5619 | -105.4261
16.5656 | -105.4246
16.5693 | -105.4222
16.5731 | -105.4174
16.5768 | -105.4135
16.5805 | -105.4091
16.5842 | -105.4043
16.588 | -105.401
16.5917 | -105.397
16.5954 | -105.3887
16.5991 | -105.3769
16.6029 | -105.3645
16.6066 | -105.3511
16.6103 | -105.3357
16.614 | -105.3172
16.6178 | -105.3004
16.6215 | -105.2802
16.6252 | -105.2589
16.6289 | -105.2386
16.6327 | -105.2172
16.6364 | -105.1967
16.6401 | -105.1753
16.6438 | -105.1503
16.6476 | -105.1264
16.6513 | -105.1004
16.655 | -105.0752
16.6587 | -105.0504
16.6625 | -105.0254
16.6662 | -105.0014
16.6699 | -104.979
16.6737 | -104.9563
16.6774 | -104.9328
16.6811 | -104.9069
16.6848 | -104.8814
16.6886 | -104.8593
16.6923 | -104.8412
16.696 | -104.8263
16.6997 | -104.8098
16.7035 | -104.7921
16.7072 | -104.7775
16.7109 | -104.7626
16.7146 | -104.7495
16.7184 | -104.7376
16.7221 | -104.7266
16.7258 | -104.7184
16.7295 | -104.7096
16.7333 | -104.701
16.737 | -104.6948
16.7407 | -104.6875
16.7444 | -104.6811
16.7482 | -104.6784
16.7519 | -104.6767
16.7556 | -104.6772
16.7594 | -104.681
16.7631 | -104.6844
16.7668 | -104.6845
16.7705 | -104.6848
16.7743 | -104.689
16.778 | -104.6976
16.7817 | -104.7075
16.7854 | -104.7187
16.7892 | -104.7249
16.7929 | -104.7312
16.7966 | -104.741
16.8003 | -104.7496
16.8041 | -104.7568
16.8078 | -104.7664
16.8115 | -104.7782
16.8152 | -104.7899
16.819 | -104.7995
16.8227 | -104.8065
16.8264 | -104.8103
16.8301 | -104.813
16.8339 | -104.8129
16.8376 | -104.8142
16.8413 | -104.8126
16.845 | -104.81
16.8488 | -104.807
16.8525 | -104.802
16.8562 | -104.7899
16.86 | -104.7758
16.8637 | -104.7614
16.8674 | -104.7424
16.8711 | -104.7236
16.8749 | -104.7067
16.8786 | -104.6874
16.8823 | -104.6653
16.886 | -104.6421
16.8898 | -104.6213
16.8935 | -104.5983
16.8972 | -104.5767
16.9009 | -104.5567
16.9047 | -104.5368
16.9084 | -104.5149
16.9121 | -104.4926
16.9158 | -104.4703
16.9196 | -104.4513
16.9233 | -104.4385
16.927 | -104.4259
16.9307 | -104.4126
16.9345 | -104.3989
16.9382 | -104.3875
16.9419 | -104.3784
16.9456 | -104.3715
16.9494 | -104.3642
16.9531 | -104.3611
16.9568 | -104.3594
16.9606 | -104.3589
16.9643 | -104.3607
16.968 | -104.3614
16.9717 | -104.3615
16.9755 | -104.3625
16.9792 | -104.3638
16.9829 | -104.3684
16.9866 | -104.3758
16.9904 | -104.3825
16.9941 | -104.3866
16.9978 | -104.3877
17.0015 | -104.3927
17.0053 | -104.4016
17.009 | -104.4126
17.0127 | -104.4232
17.0164 | -104.4296
17.0202 | -104.4311
17.0239 | -104.437
17.0276 | -104.446
17.0313 | -104.458
17.0351 | -104.4695
17.0388 | -104.4765
17.0425 | -104.4804
17.0462 | -104.4876
17.05 | -104.4932
17.0537 | -104.4968
17.0574 | -104.502
17.0612 | -104.509
17.0649 | -104.5149
17.0686 | -104.522
17.0723 | -104.53
17.0761 | -104.5408
17.0798 | -104.5481
17.0835 | -104.5528
17.0872 | -104.5561
17.091 | -104.5608
17.0947 | -104.5664
17.0984 | -104.576
17.1021 | -104.5849
17.1059 | -104.5894
17.1096 | -104.5962
17.1133 | -104.6022
17.117 | -104.6076
17.1208 | -104.6156
17.1245 | -104.6236
17.1282 | -104.6292
17.1319 | -104.6352
17.1357 | -104.6454
17.1394 | -104.6541
17.1431 | -104.6621
17.1468 | -104.6719
17.1506 | -104.6818
17.1543 | -104.6919
17.158 | -104.7018
17.1618 | -104.7113
17.1655 | -104.7182
17.1692 | -104.7252
17.1729 | -104.7323
17.1767 | -104.7392
17.1804 | -104.745
17.1841 | -104.7492
17.1878 | -104.7531
17.1916 | -104.7578
17.1953 | -104.7606
17.199 | -104.7592
17.2027 | -104.7618
17.2065 | -104.7639
17.2102 | -104.7648
17.2139 | -104.768
17.2176 | -104.7729
17.2214 | -104.7763
17.2251 | -104.7777
17.2288 | -104.774
17.2325 | -104.7707
17.2363 | -104.77
17.24 | -104.7736
17.2437 | -104.7802
17.2475 | -104.7816
17.2512 | -104.782
17.2549 | -104.7835
17.2586 | -104.7884
17.2624 | -104.7947
17.2661 | -104.8003
17.2698 | -104.8031
17.2735 | -104.8041
17.2773 | -104.8059
17.281 | -104.8085
17.2847 | -104.8139
17.2884 | -104.8231
17.2922 | -104.8289
17.2959 | -104.8344
17.2996 | -104.8401
17.3033 | -104.8438
17.3071 | -104.8481
17.3108 | -104.8538
17.3145 | -104.8593
17.3182 | -104.865
17.322 | -104.8707
17.3257 | -104.8778
17.3294 | -104.8864
17.3331 | -104.8959
17.3369 | -104.9038
17.3406 | -104.912
17.3443 | -104.9216
17.3481 | -104.9355
17.3518 | -104.9516
17.3555 | -104.9666
17.3592 | -104.9831
17.363 | -104.9987
17.3667 | -105.0157
17.3704 | -105.036
17.3741 | -105.0575
17.3779 | -105.0789
17.3816 | -105.1007
17.3853 | -105.1218
17.389 | -105.1462
17.3928 | -105.1706
17.3965 | -105.1982
17.4002 | -105.2273
17.4039 | -105.2555
17.4077 | -105.2818
17.4114 | -105.3093
17.4151 | -105.3382
17.4188 | -105.367
17.4226 | -105.3946
17.4263 | -105.4215
17.43 | -105.4499
17.4337 | -105.4773
17.4375 | -105.5018
17.4412 | -105.5222
17.4449 | -105.5435
17.4487 | -105.5606
17.4524 | -105.5772
17.4561 | -105.5921
17.4598 | -105.6036
17.4636 | -105.6124
17.4673 | -105.6194
17.471 | -105.6258
17.4747 | -105.6329
17.4785 | -105.6415
17.4822 | -105.6435
17.4859 | -105.6406
17.4896 | -105.6339
17.4934 | -105.6267
17.4971 | -105.6191
17.5008 | -105.6086
17.5045 | -105.6003
17.5083 | -105.5912
17.512 | -105.5824
17.5157 | -105.5753
17.5194 | -105.5641
17.5232 | -105.5566
17.5269 | -105.5483
17.5306 | -105.539
17.5343 | -105.5293
17.5381 | -105.5205
17.5418 | -105.5127
17.5455 | -105.5044
17.5493 | -105.4963
17.553 | -105.4869
17.5567 | -105.4794
17.5604 | -105.4744
17.5642 | -105.4685
17.5679 | -105.462
17.5716 | -105.4573
17.5753 | -105.4522
17.5791 | -105.4494
17.5828 | -105.4474
17.5865 | -105.4435
17.5902 | -105.4387
17.594 | -105.4332
17.5977 | -105.4288
17.6014 | -105.427
17.6051 | -105.4252
17.6089 | -105.4234
17.6126 | -105.4222
17.6163 | -105.4201
17.62 | -105.4182
17.6238 | -105.4181
17.6275 | -105.4196
17.6312 | -105.4196
17.6349 | -105.4207
17.6387 | -105.421
17.6424 | -105.4219
17.6461 | -105.4239
17.6499 | -105.4253
17.6536 | -105.4288
17.6573 | -105.432
17.661 | -105.4376
17.6648 | -105.4437
17.6685 | -105.4476
17.6722 | -105.4554
17.6759 | -105.4618
17.6797 | -105.4647
17.6834 | -105.47
17.6871 | -105.4751
17.6908 | -105.4842
17.6946 | -105.4941
17.6983 | -105.5017
17.702 | -105.5063
17.7057 | -105.5125
17.7095 | -105.5191
17.7132 | -105.5259
17.7169 | -105.5333
17.7206 | -105.541
17.7244 | -105.5455
17.7281 | -105.5513
17.7318 | -105.5558
17.7355 | -105.5516
17.7393 | -105.5542
17.743 | -105.5548
17.7467 | -105.5563
17.7505 | -105.5591
17.7542 | -105.5612
17.7579 | -105.5585
17.7616 | -105.5542
17.7654 | -105.55
17.7691 | -105.5452
17.7728 | -105.5374
17.7765 | -105.5258
17.7803 | -105.5157
17.784 | -105.506
17.7877 | -105.4956
17.7914 | -105.4843
17.7952 | -105.4733
17.7989 | -105.4607
17.8026 | -105.4509
17.8063 | -105.4425
17.8101 | -105.4308
17.8138 | -105.4161
17.8175 | -105.4047
17.8212 | -105.3927
17.825 | -105.3786
17.8287 | -105.3675
17.8324 | -105.3585
17.8362 | -105.3498
17.8399 | -105.343
17.8436 | -105.3371
17.8473 | -105.3286
17.8511 | -105.3201
17.8548 | -105.3128
17.8585 | -105.3081
17.8622 | -105.3028
17.866 | -105.2981
17.8697 | -105.2962
17.8734 | -105.2938
17.8771 | -105.2887
17.8809 | -105.2826
17.8846 | -105.2796
17.8883 | -105.2774
17.892 | -105.271
17.8958 | -105.2662
17.8995 | -105.2619
17.9032 | -105.2579
17.9069 | -105.2549
17.9107 | -105.252
17.9144 | -105.2447
17.9181 | -105.2401
17.9218 | -105.2381
17.9256 | -105.2354
17.9293 | -105.2343
17.933 | -105.2352
17.9368 | -105.2328
17.9405 | -105.233
17.9442 | -105.2318
17.9479 | -105.2289
17.9517 | -105.2301
17.9554 | -105.2332
17.9591 | -105.2331
17.9628 | -105.2344
17.9666 | -105.2354
17.9703 | -105.2351
17.974 | -105.236
17.9777 | -105.2414
17.9815 | -105.2454
17.9852 | -105.2505
17.9889 | -105.2586
17.9926 | -105.2665
17.9964 | -105.2748
18.0001 | -105.2854
18.0038 | -105.293
18.0075 | -105.3006
18.0113 | -105.3104
18.015 | -105.3236
18.0187 | -105.337
18.0224 | -105.347
18.0262 | -105.357
18.0299 | -105.3676
18.0336 | -105.378
18.0374 | -105.3887
18.0411 | -105.3996
18.0448 | -105.41
18.0485 | -105.4199
18.0523 | -105.4287
18.056 | -105.4384
18.0597 | -105.4448
18.0634 | -105.4501
18.0672 | -105.4545
18.0709 | -105.4596
18.0746 | -105.4619
18.0783 | -105.4642
18.0821 | -105.4679
18.0858 | -105.4679
18.0895 | -105.4631
18.0932 | -105.4566
18.097 | -105.4482
18.1007 | -105.4365
18.1044 | -105.4277
18.1081 | -105.4159
18.1119 | -105.4023
18.1156 | -105.389
18.1193 | -105.376
18.123 | -105.3628
18.1268 | -105.3479
18.1305 | -105.3334
18.1342 | -105.3187
18.138 | -105.3048
18.1417 | -105.291
18.1454 | -105.2761
18.1491 | -105.2648
18.1529 | -105.252
18.1566 | -105.2374
18.1603 | -105.223
18.164 | -105.211
18.1678 | -105.1992
18.1715 | -105.1884
18.1752 | -105.1779
18.1789 | -105.1686
18.1827 | -105.1608
18.1864 | -105.1543
18.1901 | -105.1486
18.1938 | -105.1439
18.1976 | -105.1411
18.2013 | -105.1381
18.205 | -105.1368
18.2087 | -105.1357
18.2125 | -105.1355
18.2162 | -105.1358
18.2199 | -105.1382
18.2237 | -105.1422
18.2274 | -105.1485
18.2311 | -105.155
18.2348 | -105.1608
18.2386 | -105.166
18.2423 | -105.1713
18.246 | -105.1758
18.2497 | -105.1807
18.2535 | -105.1844
18.2572 | -105.1885
18.2609 | -105.1939
18.2646 | -105.1992
18.2684 | -105.2044
18.2721 | -105.2097
18.2758 | -105.213
18.2795 | -105.2161
18.2833 | -105.2187
18.287 | -105.2219
18.2907 | -105.2259
18.2944 | -105.2269
18.2982 | -105.2225
18.3019 | -105.216
18.3056 | -105.2102
18.3093 | -105.2042
18.3131 | -105.1986
18.3168 | -105.1898
18.3205 | -105.1793
18.3243 | -105.1695
18.328 | -105.1594
18.3317 | -105.1483
18.3354 | -105.1358
18.3392 | -105.1232
18.3429 | -105.1085
18.3466 | -105.095
18.3503 | -105.081
18.3541 | -105.0654
18.3578 | -105.0525
18.3615 | -105.0407
18.3652 | -105.0276
18.369 | -105.0175
18.3727 | -105.0062
18.3764 | -104.9922
18.3801 | -104.9796
18.3839 | -104.9678
18.3876 | -104.9573
18.3913 | -104.9485
18.395 | -104.9413
18.3988 | -104.933
18.4025 | -104.9246
18.4062 | -104.9193
18.4099 | -104.9165
18.4137 | -104.9154
18.4174 | -104.9157
18.4211 | -104.9185
18.4249 | -104.9207
18.4286 | -104.9244
18.4323 | -104.9317
18.436 | -104.9394
18.4398 | -104.9465
18.4435 | -104.9492
18.4472 | -104.9559
18.4509 | -104.9645
18.4547 | -104.9726
18.4584 | -104.9802
18.4621 | -104.99
18.4658 | -104.9974
18.4696 | -105.0036
18.4733 | -105.0122
18.477 | -105.0223
18.4807 | -105.0313
18.4845 | -105.0391
18.4882 | -105.0434
18.4919 | -105.0465
18.4956 | -105.051
18.4994 | -105.0563
18.5031 | -105.0625
18.5068 | -105.0666
18.5105 | -105.0705
18.5143 | -105.0749
18.518 | -105.0791
18.5217 | -105.082
18.5255 | -105.0838
18.5292 | -105.0831
18.5329 | -105.0783
18.5366 | -105.0722
18.5404 | -105.0675
18.5441 | -105.0646
18.5478 | -105.0631
18.5515 | -105.0589
18.5553 | -105.0539
18.559 | -105.0493
18.5627 | -105.0474
18.5664 | -105.0449
18.5702 | -105.0406
18.5739 | -105.0364
18.5776 | -105.0329
18.5813 | -105.0321
18.5851 | -105.0318
18.5888 | -105.0293
18.5925 | -105.0265
18.5962 | -105.0228
18.6 | -105.0199
18.6037 | -105.0168
18.6074 | -105.0143
18.6111 | -105.0128
18.6149 | -105.0141
18.6186 | -105.0164
18.6223 | -105.0181
18.6261 | -105.0212
18.6298 | -105.0218
18.6335 | -105.0215
18.6372 | -105.0236
18.641 | -105.027
18.6447 | -105.0308
18.6484 | -105.0333
18.6521 | -105.0341
18.6559 | -105.0352
18.6596 | -105.0395
18.6633 | -105.0461
18.667 | -105.0512
18.6708 | -105.0572
18.6745 | -105.0612
18.6782 | -105.0648
18.6819 | -105.0704
18.6857 | -105.0752
18.6894 | -105.0791
18.6931 | -105.0843
18.6968 | -105.0861
18.7006 | -105.0887
18.7043 | -105.0951
18.708 | -105.1054
18.7118 | -105.1129
18.7155 | -105.119
18.7192 | -105.1245
18.7229 | -105.127
18.7267 | -105.1279
18.7304 | -105.131
18.7341 | -105.1344
18.7378 | -105.1347
18.7416 | -105.1352
18.7453 | -105.1309
18.749 | -105.1273
18.7527 | -105.1277
18.7565 | -105.1275
18.7602 | -105.1249
18.7639 | -105.122
18.7676 | -105.1187
18.7714 | -105.1159
18.7751 | -105.1134
18.7788 | -105.1074
18.7825 | -105.1002
18.7863 | -105.094
18.79 | -105.0893
18.7937 | -105.0872
18.7974 | -105.0891
18.8012 | -105.0903
18.8049 | -105.0899
18.8086 | -105.09
18.8124 | -105.0952
18.8161 | -105.0952
18.8198 | -105.0962
18.8235 | -105.0981
18.8273 | -105.0999
18.831 | -105.1019
18.8347 | -105.102
18.8384 | -105.104
18.8422 | -105.1082
18.8459 | -105.1123
18.8496 | -105.1165
18.8533 | -105.1214
18.8571 | -105.1282
18.8608 | -105.1365
18.8645 | -105.144
18.8682 | -105.1485
18.872 | -105.1532
18.8757 | -105.1583
18.8794 | -105.1624
18.8831 | -105.167
18.8869 | -105.1717
18.8906 | -105.1749
18.8943 | -105.178
18.898 | -105.1835
18.9018 | -105.1867
18.9055 | -105.1895
18.9092 | -105.1927
18.913 | -105.1947
18.9167 | -105.1972
18.9204 | -105.1999
18.9241 | -105.2007
18.9279 | -105.199
18.9316 | -105.1979
18.9353 | -105.1972
18.939 | -105.1959
18.9428 | -105.1958
18.9465 | -105.1945
18.9502 | -105.1931
18.9539 | -105.1934
18.9577 | -105.1956
18.9614 | -105.1981
18.9651 | -105.2038
18.9688 | -105.2081
18.9726 | -105.2135
18.9763 | -105.2197
18.98 | -105.2245
18.9837 | -105.2288
18.9875 | -105.2343
18.9912 | -105.2404
18.9949 | -105.2501
18.9986 | -105.2619
19.0024 | -105.2729
19.0061 | -105.2805
19.0098 | -105.2876
19.0136 | -105.2981
19.0173 | -105.3111
19.021 | -105.3217
19.0247 | -105.3306
19.0285 | -105.3408
19.0322 | -105.3522
19.0359 | -105.3633
19.0396 | -105.3726
19.0434 | -105.3819
19.0471 | -105.3935
19.0508 | -105.4038
19.0545 | -105.4108
19.0583 | -105.4171
19.062 | -105.4228
19.0657 | -105.4269
19.0694 | -105.4314
19.0732 | -105.4364
19.0769 | -105.4396
19.0806 | -105.4419
19.0843 | -105.4465
19.0881 | -105.4529
19.0918 | -105.4564
19.0955 | -105.4581
19.0993 | -105.4615
19.103 | -105.4633
19.1067 | -105.4665
19.1104 | -105.4697
19.1142 | -105.4723
19.1179 | -105.4752
19.1216 | -105.4773
19.1253 | -105.4804
19.1291 | -105.4843
19.1328 | -105.487
19.1365 | -105.4917
19.1402 | -105.4998
19.144 | -105.5069
19.1477 | -105.5148
19.1514 | -105.5216
19.1551 | -105.5266
19.1589 | -105.5305
19.1626 | -105.5345
19.1663 | -105.5371
19.17 | -105.5417
19.1738 | -105.5464
19.1775 | -105.549
19.1812 | -105.551
19.1849 | -105.5546
19.1887 | -105.5584
19.1924 | -105.5611
19.1961 | -105.5616
19.1999 | -105.5601
19.2036 | -105.5567
19.2073 | -105.5535
19.211 | -105.548
19.2148 | -105.5419
19.2185 | -105.5368
19.2222 | -105.5305
19.2259 | -105.5231
19.2297 | -105.5163
19.2334 | -105.5082
19.2371 | -105.4985
19.2408 | -105.4869
19.2446 | -105.4717
19.2483 | -105.459
19.252 | -105.4481
19.2557 | -105.4389
19.2595 | -105.4312
19.2632 | -105.423
19.2669 | -105.4108
19.2706 | -105.3971
19.2744 | -105.382
19.2781 | -105.3697
19.2818 | -105.3578
19.2855 | -105.3463
19.2893 | -105.3341
19.293 | -105.3236
19.2967 | -105.3149
19.3005 | -105.3078
19.3042 | -105.3014
19.3079 | -105.2941
19.3116 | -105.288
19.3154 | -105.2833
19.3191 | -105.2806
19.3228 | -105.2766
19.3265 | -105.2741
19.3303 | -105.2742
19.334 | -105.275
19.3377 | -105.2756
19.3414 | -105.2772
19.3452 | -105.2791
19.3489 | -105.2874
19.3526 | -105.2908
19.3563 | -105.2958
19.3601 | -105.299
19.3638 | -105.3019
19.3675 | -105.306
19.3712 | -105.3104
19.375 | -105.3159
19.3787 | -105.3206
19.3824 | -105.3276
19.3861 | -105.3335
19.3899 | -105.3372
19.3936 | -105.3413
19.3973 | -105.3475
19.4011 | -105.3517
19.4048 | -105.3545
19.4085 | -105.358
19.4122 | -105.3623
19.416 | -105.3649
19.4197 | -105.3673
19.4234 | -105.3687
19.4271 | -105.3718
19.4309 | -105.3754
19.4346 | -105.3789
19.4383 | -105.3814
19.442 | -105.3841
19.4458 | -105.3853
19.4495 | -105.3864
19.4532 | -105.3867
19.4569 | -105.3864
19.4607 | -105.384
19.4644 | -105.3815
19.4681 | -105.3807
19.4718 | -105.3814
19.4756 | -105.3819
19.4793 | -105.3819
19.483 | -105.3815
19.4868 | -105.3792
19.4905 | -105.3783
19.4942 | -105.3754
19.4979 | -105.3697
19.5017 | -105.3618
19.5054 | -105.3555
19.5091 | -105.3515
19.5128 | -105.3487
19.5166 | -105.349
19.5203 | -105.3488
19.524 | -105.3458
19.5277 | -105.3439
19.5315 | -105.3412
19.5352 | -105.3379
19.5389 | -105.3356
19.5426 | -105.3332
19.5464 | -105.3313
19.5501 | -105.3312
19.5538 | -105.3314
19.5575 | -105.3304
19.5613 | -105.3265
19.565 | -105.3198
19.5687 | -105.3124
19.5724 | -105.3079
19.5762 | -105.3017
19.5799 | -105.2966
19.5836 | -105.2951
19.5874 | -105.2937
19.5911 | -105.2907
19.5948 | -105.2858
19.5985 | -105.2849
19.6023 | -105.2851
19.606 | -105.2811
19.6097 | -105.2796
19.6134 | -105.2793
19.6172 | -105.2798
19.6209 | -105.2788
19.6246 | -105.2766
19.6283 | -105.2765
19.6321 | -105.2778
19.6358 | -105.2788
19.6395 | -105.2805
19.6432 | -105.2824
19.647 | -105.2859
19.6507 | -105.288
19.6544 | -105.289
19.6581 | -105.292
19.6619 | -105.2908
19.6656 | -105.2915
19.6693 | -105.2948
19.673 | -105.2992
19.6768 | -105.303
19.6805 | -105.3076
19.6842 | -105.3087
19.688 | -105.3062
19.6917 | -105.3069
19.6954 | -105.308
19.6991 | -105.3083
19.7029 | -105.3095
19.7066 | -105.3088
19.7103 | -105.3069
19.714 | -105.3067
19.7178 | -105.3073
19.7215 | -105.307
19.7252 | -105.3034
19.7289 | -105.2989
19.7327 | -105.2943
19.7364 | -105.2922
19.7401 | -105.2905
19.7438 | -105.2855
19.7476 | -105.2789
19.7513 | -105.271
19.755 | -105.2631
19.7587 | -105.2551
19.7625 | -105.2463
19.7662 | -105.2397
19.7699 | -105.2316
19.7736 | -105.2243
19.7774 | -105.2175
19.7811 | -105.2106
19.7848 | -105.201
19.7886 | -105.1873
19.7923 | -105.1747
19.796 | -105.1629
19.7997 | -105.1544
19.8035 | -105.1459
19.8072 | -105.1373
19.8109 | -105.1295
19.8146 | -105.1227
19.8184 | -105.1163
19.8221 | -105.1107
19.8258 | -105.106
19.8295 | -105.1031
19.8333 | -105.1011
19.837 | -105.1006
19.8407 | -105.1016
19.8444 | -105.1038
19.8482 | -105.1033
19.8519 | -105.1007
19.8556 | -105.0958
19.8593 | -105.0882
19.8631 | -105.0806
19.8668 | -105.0756
19.8705 | -105.072
19.8742 | -105.0689
19.878 | -105.0647
19.8817 | -105.0598
19.8854 | -105.0494
19.8892 | -105.0418
19.8929 | -105.0324
19.8966 | -105.0231
19.9003 | -105.0126
19.9041 | -105.002
19.9078 | -104.9911
19.9115 | -104.9775
19.9152 | -104.9636
19.919 | -104.9478
19.9227 | -104.931
19.9264 | -104.9132
19.9301 | -104.8927
19.9339 | -104.8726
19.9376 | -104.8529
19.9413 | -104.8326
19.945 | -104.8139
19.9488 | -104.7929
19.9525 | -104.7692
19.9562 | -104.7412
19.9599 | -104.7134
19.9637 | -104.6878
19.9674 | -104.665
19.9711 | -104.6428
19.9748 | -104.6176
19.9786 | -104.5936
19.9823 | -104.5716
19.986 | -104.5475
19.9898 | -104.523
19.9935 | -104.5001
19.9972 | -104.476
20.0009 | -104.4486
20.0047 | -104.4178
20.0084 | -104.3932
20.0121 | -104.3744
20.0158 | -104.3566
20.0196 | -104.3357
20.0233 | -104.3147
20.027 | -104.2966
20.0307 | -104.2795
20.0345 | -104.2602
20.0382 | -104.2408
20.0419 | -104.2235
20.0456 | -104.2058
20.0494 | -104.1859
20.0531 | -104.1634
20.0568 | -104.1431
20.0605 | -104.1265
20.0643 | -104.111
20.068 | -104.0954
20.0717 | -104.0815
20.0755 | -104.0647
20.0792 | -104.0428
20.0829 | -104.021
20.0866 | -104.0021
20.0904 | -103.9837
20.0941 | -103.9638
20.0978 | -103.9432
20.1015 | -103.9238
20.1053 | -103.9064
20.109 | -103.8901
20.1127 | -103.8732
20.1164 | -103.8566
20.1202 | -103.8421
20.1239 | -103.8276
20.1276 | -103.8126
20.1313 | -103.7988
20.1351 | -103.7868
20.1388 | -103.779
20.1425 | -103.7699
20.1462 | -103.7564
20.15 | -103.7387
20.1537 | -103.724
20.1574 | -103.7117
20.1611 | -103.7011
20.1649 | -103.691
20.1686 | -103.6805
20.1723 | -103.6716
20.1761 | -103.6644
20.1798 | -103.6556
20.1835 | -103.646
20.1872 | -103.6366
20.191 | -103.6269
20.1947 | -103.617
20.1984 | -103.6082
20.2021 | -103.5974
20.2059 | -103.5868
20.2096 | -103.5773
20.2133 | -103.567
20.217 | -103.5578
20.2208 | -103.5494
20.2245 | -103.5389
20.2282 | -103.528
20.2319 | -103.5175
20.2357 | -103.5079
20.2394 | -103.4981
20.2431 | -103.4916
20.2468 | -103.4827
20.2506 | -103.4733
20.2543 | -103.4666
20.258 | -103.4606
20.2617 | -103.4514
20.2655 | -103.4425
20.2692 | -103.4366
20.2729 | -103.4337
20.2767 | -103.4322
20.2804 | -103.4291
20.2841 | -103.4258
20.2878 | -103.4277
20.2916 | -103.4321
20.2953 | -103.428
20.299 | -103.4175
20.3027 | -103.4105
20.3065 | -103.4078
20.3102 | -103.4056
20.3139 | -103.4043
20.3176 | -103.4042
20.3214 | -103.4061
20.3251 | -103.4089
20.3288 | -103.4073
20.3325 | -103.4015
20.3363 | -103.3993
20.34 | -103.3973
20.3437 | -103.3888
20.3474 | -103.3778
20.3512 | -103.3681
20.3549 | -103.3564
20.3586 | -103.346
20.3623 | -103.337
20.3661 | -103.3291
20.3698 | -103.3187
20.3735 | -103.3053
20.3773 | -103.2908
20.381 | -103.2784
20.3847 | -103.2648
20.3884 | -103.2503
20.3922 | -103.2349
20.3959 | -103.218
20.3996 | -103.2028
20.4033 | -103.1882
20.4071 | -103.173
20.4108 | -103.1585
20.4145 | -103.1454
20.4182 | -103.1335
20.422 | -103.1206
20.4257 | -103.1135
20.4294 | -103.1052
20.4331 | -103.101
20.4369 | -103.0999
20.4406 | -103.1005
20.4443 | -103.098
20.448 | -103.0965
20.4518 | -103.0953
20.4555 | -103.0956
20.4592 | -103.0983
20.4629 | -103.1019
20.4667 | -103.1078
20.4704 | -103.1167
20.4741 | -103.1268
20.4779 | -103.1389
20.4816 | -103.154
20.4853 | -103.1689
20.489 | -103.1827
20.4928 | -103.1993
20.4965 | -103.2158
20.5002 | -103.2313
20.5039 | -103.2482
20.5077 | -103.263
20.5114 | -103.2761
20.5151 | -103.2886
20.5188 | -103.3005
20.5226 | -103.313
20.5263 | -103.326
20.53 | -103.3406
20.5337 | -103.3538
20.5375 | -103.3688
20.5412 | -103.3807
20.5449 | -103.3894
20.5486 | -103.3971
20.5524 | -103.4062
20.5561 | -103.4137
20.5598 | -103.4154
20.5636 | -103.4141
20.5673 | -103.414
20.571 | -103.4163
20.5747 | -103.4194
20.5785 | -103.4199
20.5822 | -103.4187
20.5859 | -103.4213
20.5896 | -103.4231
20.5934 | -103.419
20.5971 | -103.4121
20.6008 | -103.4073
20.6045 | -103.4006
20.6083 | -103.3914
20.612 | -103.3859
20.6157 | -103.3808
20.6194 | -103.3768
20.6232 | -103.3743
20.6269 | -103.3664
20.6306 | -103.357
20.6343 | -103.3506
20.6381 | -103.3456
20.6418 | -103.3387
20.6455 | -103.331
20.6492 | -103.3224
20.653 | -103.3156
20.6567 | -103.3109
20.6604 | -103.3054
20.6642 | -103.2997
20.6679 | -103.2926
20.6716 | -103.2822
20.6753 | -103.2758
20.6791 | -103.2703
20.6828 | -103.2648
20.6865 | -103.2601
20.6902 | -103.2557
20.694 | -103.251
20.6977 | -103.2452
20.7014 | -103.2408
20.7051 | -103.2355
20.7089 | -103.2261
20.7126 | -103.2187
20.7163 | -103.2181
20.72 | -103.2149
20.7238 | -103.2056
20.7275 | -103.197
20.7312 | -103.1833
20.7349 | -103.1678
20.7387 | -103.1538
20.7424 | -103.1376
20.7461 | -103.124
20.7498 | -103.1117
20.7536 | -103.0986
20.7573 | -103.0832
20.761 | -103.0664
20.7648 | -103.0509
20.7685 | -103.0358
20.7722 | -103.0177
20.7759 | -102.9963
20.7797 | -102.9755
20.7834 | -102.9562
20.7871 | -102.935
20.7908 | -102.9137
20.7946 | -102.8923
20.7983 | -102.8692
20.802 | -102.8481
20.8057 | -102.8271
20.8095 | -102.8072
20.8132 | -102.7869
20.8169 | -102.7626
20.8206 | -102.7377
20.8244 | -102.7134
20.8281 | -102.6897
20.8318 | -102.6651
20.8355 | -102.6394
20.8393 | -102.6138
20.843 | -102.5919
20.8467 | -102.5745
20.8504 | -102.5584
20.8542 | -102.5388
20.8579 | -102.5238
20.8616 | -102.5123
20.8654 | -102.5014
20.8691 | -102.4925
20.8728 | -102.4858
20.8765 | -102.4768
20.8803 | -102.461
20.884 | -102.4373
20.8877 | -102.4238
20.8914 | -102.4194
20.8952 | -102.4158
20.8989 | -102.4128
20.9026 | -102.4118
20.9063 | -102.4103
20.9101 | -102.4087
20.9138 | -102.41
20.9175 | -102.4122
20.9212 | -102.4139
20.925 | -102.4173
20.9287 | -102.4199
20.9324 | -102.4195
20.9361 | -102.4181
20.9399 | -102.4172
20.9436 | -102.4154
20.9473 | -102.4118
20.951 | -102.4067
20.9548 | -102.4008
20.9585 | -102.3944
20.9622 | -102.3846
20.966 | -102.3806
20.9697 | -102.3771
20.9734 | -102.3704
20.9771 | -102.3642
20.9809 | -102.3574
20.9846 | -102.3516
20.9883 | -102.3473
20.992 | -102.3433
20.9958 | -102.3414
20.9995 | -102.3382
21.0032 | -102.338
21.0069 | -102.3431
21.0107 | -102.3467
21.0144 | -102.3484
21.0181 | -102.3528
21.0218 | -102.3595
21.0256 | -102.3674
21.0293 | -102.3804
21.033 | -102.3973
21.0367 | -102.4139
21.0405 | -102.4303
21.0442 | -102.4454
21.0479 | -102.4621
21.0517 | -102.4797
21.0554 | -102.4996
21.0591 | -102.5196
21.0628 | -102.5383
21.0666 | -102.5565
21.0703 | -102.5764
21.074 | -102.5956
21.0777 | -102.617
21.0815 | -102.6385
21.0852 | -102.658
21.0889 | -102.6783
21.0926 | -102.6972
21.0964 | -102.7158
21.1001 | -102.734
21.1038 | -102.7496
21.1075 | -102.7635
21.1113 | -102.7781
21.115 | -102.7949
21.1187 | -102.8112
21.1224 | -102.8257
21.1262 | -102.8386
21.1299 | -102.8519
21.1336 | -102.8678
21.1373 | -102.8832
21.1411 | -102.8947
21.1448 | -102.9003
21.1485 | -102.9065
21.1523 | -102.9204
21.156 | -102.9295
21.1597 | -102.9331
21.1634 | -102.9383
21.1672 | -102.9361
21.1709 | -102.9263
21.1746 | -102.9212
21.1783 | -102.9279
21.1821 | -102.9451
21.1858 | -102.9487
21.1895 | -102.9419
21.1932 | -102.9348
21.197 | -102.9258
21.2007 | -102.9187
21.2044 | -102.9132
21.2081 | -102.9081
21.2119 | -102.903
21.2156 | -102.8992
21.2193 | -102.8976
21.223 | -102.8946
21.2268 | -102.8907
21.2305 | -102.886
21.2342 | -102.8819
21.2379 | -102.8802
21.2417 | -102.8769
21.2454 | -102.8723
21.2491 | -102.8653
21.2529 | -102.859
21.2566 | -102.856
21.2603 | -102.8538
21.264 | -102.8505
21.2678 | -102.8465
21.2715 | -102.8435
21.2752 | -102.8393
21.2789 | -102.8344
21.2827 | -102.83
21.2864 | -102.8253
21.2901 | -102.8201
21.2938 | -102.8156
21.2976 | -102.8114
21.3013 | -102.8105
21.305 | -102.8089
21.3087 | -102.8071
21.3125 | -102.8068
21.3162 | -102.8034
21.3199 | -102.7994
21.3236 | -102.7998
21.3274 | -102.8027
21.3311 | -102.8078
21.3348 | -102.811
21.3386 | -102.8144
21.3423 | -102.8203
21.346 | -102.8281
21.3497 | -102.8348
21.3535 | -102.8431
21.3572 | -102.8511
21.3609 | -102.8592
21.3646 | -102.8692
21.3684 | -102.8816
21.3721 | -102.8948
21.3758 | -102.9084
21.3795 | -102.9172
21.3833 | -102.9257
21.387 | -102.9371
21.3907 | -102.9464
21.3944 | -102.9576
21.3982 | -102.9672
21.4019 | -102.9722
21.4056 | -102.977
21.4093 | -102.9817
21.4131 | -102.9859
21.4168 | -102.9866
21.4205 | -102.9862
21.4242 | -102.9875
21.428 | -102.9903
21.4317 | -102.9936
21.4354 | -102.9953
21.4392 | -102.9971
21.4429 | -102.9991
21.4466 | -103.003
21.4503 | -103.0072
21.4541 | -103.0118
21.4578 | -103.0183
21.4615 | -103.0247
21.4652 | -103.028
21.469 | -103.0379
21.4727 | -103.0554
21.4764 | -103.0719
21.4801 | -103.0825
21.4839 | -103.0897
21.4876 | -103.0984
21.4913 | -103.1067
21.495 | -103.1137
21.4988 | -103.1227
21.5025 | -103.1323
21.5062 | -103.1393
21.5099 | -103.1435
21.5137 | -103.1493
21.5174 | -103.1557
21.5211 | -103.1633
21.5248 | -103.1705
21.5286 | -103.1783
21.5323 | -103.1882
21.536 | -103.1984
21.5398 | -103.2076
21.5435 | -103.2135
21.5472 | -103.2205
21.5509 | -103.2311
21.5547 | -103.2435
21.5584 | -103.2544
21.5621 | -103.2612
21.5658 | -103.2639
21.5696 | -103.2664
21.5733 | -103.2689
21.577 | -103.2716
21.5807 | -103.2768
21.5845 | -103.282
21.5882 | -103.2859
21.5919 | -103.2899
21.5956 | -103.2934
21.5994 | -103.2957
21.6031 | -103.2988
21.6068 | -103.2998
21.6105 | -103.2995
21.6143 | -103.2994
21.618 | -103.2983
21.6217 | -103.3006
21.6254 | -103.3023
21.6292 | -103.3014
21.6329 | -103.2983
21.6366 | -103.2934
21.6404 | -103.2881
21.6441 | -103.2828
21.6478 | -103.2756
21.6515 | -103.2684
21.6553 | -103.2634
21.659 | -103.2582
21.6627 | -103.2542
21.6664 | -103.2512
21.6702 | -103.2457
21.6739 | -103.2395
21.6776 | -103.2349
21.6813 | -103.2301
21.6851 | -103.2253
21.6888 | -103.2165
21.6925 | -103.2071
21.6962 | -103.2016
21.7 | -103.196
21.7037 | -103.1922
21.7074 | -103.1917
21.7111 | -103.1917
21.7149 | -103.1912
21.7186 | -103.1896
21.7223 | -103.1913
21.726 | -103.1921
21.7298 | -103.1919
21.7335 | -103.1919
21.7372 | -103.1924
21.741 | -103.1947
21.7447 | -103.1984
21.7484 | -103.2038
21.7521 | -103.2063
21.7559 | -103.1997
21.7596 | -103.1864
21.7633 | -103.1794
21.767 | -103.181
21.7708 | -103.1858
21.7745 | -103.1891
21.7782 | -103.1883
21.7819 | -103.189
21.7857 | -103.1902
21.7894 | -103.1912
21.7931 | -103.1928
21.7968 | -103.1932
21.8006 | -103.1921
21.8043 | -103.1925
21.808 | -103.1911
21.8117 | -103.1885
21.8155 | -103.1875
21.8192 | -103.185
21.8229 | -103.1817
21.8267 | -103.1759
21.8304 | -103.168
21.8341 | -103.1606
21.8378 | -103.1514
21.8416 | -103.1393
21.8453 | -103.1269
21.849 | -103.1154
21.8527 | -103.1051
21.8565 | -103.093
21.8602 | -103.0783
21.8639 | -103.0632
21.8676 | -103.0494
21.8714 | -103.0341
21.8751 | -103.0222
21.8788 | -103.0087
21.8825 | -102.9936
21.8863 | -102.9764
21.89 | -102.9584
21.8937 | -102.9406
21.8974 | -102.9244
21.9012 | -102.9065
21.9049 | -102.8803
21.9086 | -102.8581
21.9123 | -102.8386
21.9161 | -102.8223
21.9198 | -102.8077
21.9235 | -102.7928
21.9273 | -102.7791
21.931 | -102.7638
21.9347 | -102.7497
21.9384 | -102.7351
21.9422 | -102.7215
21.9459 | -102.7099
21.9496 | -102.6994
21.9533 | -102.689
21.9571 | -102.6772
21.9608 | -102.6636
21.9645 | -102.6497
21.9682 | -102.6371
21.972 | -102.626
21.9757 | -102.6149
21.9794 | -102.6051
21.9831 | -102.5952
21.9869 | -102.5851
21.9906 | -102.5769
21.9943 | -102.5712
21.998 | -102.5666
22.0018 | -102.5626
22.0055 | -102.5561
22.0092 | -102.5479
22.0129 | -102.5414
22.0167 | -102.5388
22.0204 | -102.5379
22.0241 | -102.5383
22.0279 | -102.5376
22.0316 | -102.5376
22.0353 | -102.5405
22.039 | -102.5557
22.0428 | -102.563
22.0465 | -102.5713
22.0502 | -102.5778
22.0539 | -102.5843
22.0577 | -102.593
22.0614 | -102.6031
22.0651 | -102.6151
22.0688 | -102.6345
22.0726 | -102.6582
22.0763 | -102.6748
22.08 | -102.6871
22.0837 | -102.7005
22.0875 | -102.7138
22.0912 | -102.7271
22.0949 | -102.741
22.0986 | -102.7558
22.1024 | -102.7724
22.1061 | -102.7878
22.1098 | -102.8038
22.1135 | -102.8214
22.1173 | -102.8391
22.121 | -102.8556
22.1247 | -102.8704
22.1285 | -102.8825
22.1322 | -102.8954
22.1359 | -102.9064
22.1396 | -102.9176
22.1434 | -102.9274
22.1471 | -102.9337
22.1508 | -102.9424
22.1545 | -102.9512
22.1583 | -102.9561
22.162 | -102.9603
22.1657 | -102.9622
22.1694 | -102.9637
22.1732 | -102.9614
22.1769 | -102.9586
22.1806 | -102.9545
22.1843 | -102.9522
22.1881 | -102.9467
22.1918 | -102.9361
22.1955 | -102.921
22.1992 | -102.9085
22.203 | -102.8979
22.2067 | -102.8866
22.2104 | -102.8737
22.2141 | -102.8604
22.2179 | -102.8525
22.2216 | -102.8401
22.2253 | -102.8223
22.2291 | -102.8011
22.2328 | -102.7797
22.2365 | -102.758
22.2402 | -102.7348
22.244 | -102.7133
22.2477 | -102.6912
22.2514 | -102.6699
22.2551 | -102.6472
22.2589 | -102.6246
22.2626 | -102.6017
22.2663 | -102.5786
22.27 | -102.5561
22.2738 | -102.5354
22.2775 | -102.5141
22.2812 | -102.4935
22.2849 | -102.4737
22.2887 | -102.4544
22.2924 | -102.4346
22.2961 | -102.4157
22.2998 | -102.398
22.3036 | -102.3817
22.3073 | -102.3674
22.311 | -102.3556
22.3148 | -102.346
22.3185 | -102.3367
22.3222 | -102.3285
22.3259 | -102.3224
22.3297 | -102.3167
22.3334 | -102.3122
22.3371 | -102.3092
22.3408 | -102.3083
22.3446 | -102.308
22.3483 | -102.3086
22.352 | -102.3115
22.3557 | -102.3159
22.3595 | -102.3221
22.3632 | -102.3304
22.3669 | -102.3415
22.3706 | -102.3548
22.3744 | -102.3676
22.3781 | -102.3805
22.3818 | -102.3952
22.3855 | -102.4095
22.3893 | -102.4245
22.393 | -102.44
22.3967 | -102.4523
22.4004 | -102.4641
22.4042 | -102.4769
22.4079 | -102.4922
22.4116 | -102.5053
22.4154 | -102.5176
22.4191 | -102.529
22.4228 | -102.5389
22.4265 | -102.5494
22.4303 | -102.5612
22.434 | -102.5724
22.4377 | -102.5856
22.4414 | -102.598
22.4452 | -102.6094
22.4489 | -102.6181
22.4526 | -102.6267
22.4563 | -102.6353
22.4601 | -102.6443
22.4638 | -102.6549
22.4675 | -102.6618
22.4712 | -102.6684
22.475 | -102.6781
22.4787 | -102.6873
22.4824 | -102.6957
22.4861 | -102.7021
22.4899 | -102.7064
22.4936 | -102.7088
22.4973 | -102.7141
22.501 | -102.7174
22.5048 | -102.7208
22.5085 | -102.7262
22.5122 | -102.732
22.516 | -102.7353
22.5197 | -102.7357
22.5234 | -102.7335
22.5271 | -102.7301
22.5309 | -102.7258
22.5346 | -102.7221
22.5383 | -102.7198
22.542 | -102.7161
22.5458 | -102.7103
22.5495 | -102.7001
22.5532 | -102.6896
22.5569 | -102.6809
22.5607 | -102.6723
22.5644 | -102.6645
22.5681 | -102.6534
22.5718 | -102.6408
22.5756 | -102.6227
22.5793 | -102.6091
22.583 | -102.5958
22.5867 | -102.5838
22.5905 | -102.5712
22.5942 | -102.557
22.5979 | -102.5434
22.6016 | -102.5295
22.6054 | -102.5153
22.6091 | -102.5031
22.6128 | -102.491
22.6166 | -102.4782
22.6203 | -102.4675
22.624 | -102.4556
22.6277 | -102.4399
22.6315 | -102.424
22.6352 | -102.4068
22.6389 | -102.3931
22.6426 | -102.3829
22.6464 | -102.3717
22.6501 | -102.3594
22.6538 | -102.3487
22.6575 | -102.3397
22.6613 | -102.3331
22.665 | -102.3258
22.6687 | -102.3164
22.6724 | -102.3077
22.6762 | -102.3
22.6799 | -102.2926
22.6836 | -102.286
22.6873 | -102.2807
22.6911 | -102.2759
22.6948 | -102.2723
22.6985 | -102.2692
22.7022 | -102.2681
22.706 | -102.2682
22.7097 | -102.2679
22.7134 | -102.2692
22.7172 | -102.2722
22.7209 | -102.2726
22.7246 | -102.2721
22.7283 | -102.2727
22.7321 | -102.2759
22.7358 | -102.2783
22.7395 | -102.2813
22.7432 | -102.2861
22.747 | -102.2886
22.7507 | -102.2896
22.7544 | -102.2912
22.7581 | -102.2923
22.7619 | -102.295
22.7656 | -102.2998
22.7693 | -102.3067
22.773 | -102.3131
22.7768 | -102.3137
22.7805 | -102.3092
22.7842 | -102.3034
22.7879 | -102.3006
22.7917 | -102.2971
22.7954 | -102.2937
22.7991 | -102.2912
22.8028 | -102.2889
22.8066 | -102.2882
22.8103 | -102.2898
22.814 | -102.2886
22.8178 | -102.282
22.8215 | -102.2767
22.8252 | -102.2711
22.8289 | -102.2651
22.8327 | -102.2593
22.8364 | -102.251
22.8401 | -102.2377
22.8438 | -102.2258
22.8476 | -102.2178
22.8513 | -102.2098
22.855 | -102.1989
22.8587 | -102.1849
22.8625 | -102.1722
22.8662 | -102.1594
22.8699 | -102.1476
22.8736 | -102.1354
22.8774 | -102.1237
22.8811 | -102.1116
22.8848 | -102.1004
22.8885 | -102.0905
22.8923 | -102.0818
22.896 | -102.0723
22.8997 | -102.0648
22.9035 | -102.0572
22.9072 | -102.0491
22.9109 | -102.0387
22.9146 | -102.0292
22.9184 | -102.0191
22.9221 | -102.0114
22.9258 | -102.0066
22.9295 | -102.0032
22.9333 | -102.0005
22.937 | -101.9997
22.9407 | -102.0018
22.9444 | -102.0034
22.9482 | -102.007
22.9519 | -102.0121
22.9556 | -102.0187
22.9593 | -102.0276
22.9631 | -102.0354
22.9668 | -102.0419
22.9705 | -102.0473
22.9742 | -102.0529
22.978 | -102.0609
22.9817 | -102.0688
22.9854 | -102.0775
22.9891 | -102.0874
22.9929 | -102.0993
22.9966 | -102.1105
23.0003 | -102.1203
23.0041 | -102.1275
23.0078 | -102.1359
23.0115 | -102.145
23.0152 | -102.1513
23.019 | -102.1572
23.0227 | -102.1644
23.0264 | -102.171
23.0301 | -102.1763
23.0339 | -102.1791
23.0376 | -102.1796
23.0413 | -102.1796
23.045 | -102.1838
23.0488 | -102.1889
23.0525 | -102.1902
23.0562 | -102.1876
23.0599 | -102.185
23.0637 | -102.1808
23.0674 | -102.1781
23.0711 | -102.177
23.0748 | -102.1735
23.0786 | -102.1674
23.0823 | -102.1624
23.086 | -102.1573
23.0897 | -102.1517
23.0935 | -102.1454
23.0972 | -102.1388
23.1009 | -102.137
23.1047 | -102.1394
23.1084 | -102.14
23.1121 | -102.1356
23.1158 | -102.1317
23.1196 | -102.1279
23.1233 | -102.1222
23.127 | -102.1157
23.1307 | -102.112
23.1345 | -102.1078
23.1382 | -102.1042
23.1419 | -102.1026
23.1456 | -102.1026
23.1494 | -102.1016
23.1531 | -102.0998
23.1568 | -102.0991
23.1605 | -102.099
23.1643 | -102.1006
23.168 | -102.1059
23.1717 | -102.1096
23.1754 | -102.112
23.1792 | -102.1154
23.1829 | -102.1215
23.1866 | -102.1304
23.1903 | -102.1404
23.1941 | -102.1525
23.1978 | -102.1663
23.2015 | -102.1805
23.2053 | -102.1933
23.209 | -102.2058
23.2127 | -102.2164
23.2164 | -102.2269
23.2202 | -102.238
23.2239 | -102.2484
23.2276 | -102.262
23.2313 | -102.2776
23.2351 | -102.2948
23.2388 | -102.3127
23.2425 | -102.3312
23.2462 | -102.3519
23.25 | -102.3693
23.2537 | -102.3837
23.2574 | -102.3965
23.2611 | -102.4095
23.2649 | -102.4223
23.2686 | -102.4342
23.2723 | -102.4446
23.276 | -102.4533
23.2798 | -102.4631
23.2835 | -102.4727
23.2872 | -102.481
23.291 | -102.4896
23.2947 | -102.4979
23.2984 | -102.5063
23.3021 | -102.515
23.3059 | -102.5191
23.3096 | -102.5221
23.3133 | -102.5258
23.317 | -102.5276
23.3208 | -102.5262
23.3245 | -102.5257
23.3282 | -102.5251
23.3319 | -102.5282
23.3357 | -102.5284
23.3394 | -102.5269
23.3431 | -102.5285
23.3468 | -102.5286
23.3506 | -102.5249
23.3543 | -102.5218
23.358 | -102.5187
23.3617 | -102.5154
23.3655 | -102.5122
23.3692 | -102.5086
23.3729 | -102.5033
23.3766 | -102.4977
23.3804 | -102.4916
23.3841 | -102.4854
23.3878 | -102.4791
23.3916 | -102.4715
23.3953 | -102.4619
23.399 | -102.4547
23.4027 | -102.4483
23.4065 | -102.4403
23.4102 | -102.4314
23.4139 | -102.4228
23.4176 | -102.4152
23.4214 | -102.4089
23.4251 | -102.4024
23.4288 | -102.3957
23.4325 | -102.3886
23.4363 | -102.381
23.44 | -102.3755
23.4437 | -102.3706
23.4474 | -102.366
23.4512 | -102.3625
23.4549 | -102.3602
23.4586 | -102.3578
23.4623 | -102.3552
23.4661 | -102.352
23.4698 | -102.3509
23.4735 | -102.3521
23.4772 | -102.3527
23.481 | -102.353
23.4847 | -102.3554
23.4884 | -102.3604
23.4922 | -102.3662
23.4959 | -102.3712
23.4996 | -102.3754
23.5033 | -102.3814
23.5071 | -102.3851
23.5108 | -102.3901
23.5145 | -102.3954
23.5182 | -102.3999
23.522 | -102.4061
23.5257 | -102.4123
23.5294 | -102.4179
23.5331 | -102.4254
23.5369 | -102.4321
23.5406 | -102.4384
23.5443 | -102.4499
23.548 | -102.4636
23.5518 | -102.4751
23.5555 | -102.4874
23.5592 | -102.4975
23.5629 | -102.5062
23.5667 | -102.5156
23.5704 | -102.5272
23.5741 | -102.538
23.5778 | -102.5455
23.5816 | -102.5532
23.5853 | -102.5619
23.589 | -102.5714
23.5928 | -102.5817
23.5965 | -102.5916
23.6002 | -102.6014
23.6039 | -102.6114
23.6077 | -102.6223
23.6114 | -102.6319
23.6151 | -102.637
23.6188 | -102.6387
23.6226 | -102.6392
23.6263 | -102.6415
23.63 | -102.6447
23.6337 | -102.6477
23.6375 | -102.6511
23.6412 | -102.6525
23.6449 | -102.6504
23.6486 | -102.6409
23.6524 | -102.6383
23.6561 | -102.6389
23.6598 | -102.6373
23.6635 | -102.635
23.6673 | -102.6312
23.671 | -102.6285
23.6747 | -102.6262
23.6785 | -102.622
23.6822 | -102.6183
23.6859 | -102.6162
23.6896 | -102.6118
23.6934 | -102.6089
23.6971 | -102.608
23.7008 | -102.6057
23.7045 | -102.6033
23.7083 | -102.5996
23.712 | -102.5956
23.7157 | -102.5936
23.7194 | -102.5896
23.7232 | -102.5832
23.7269 | -102.5798
23.7306 | -102.5787
23.7343 | -102.5755
23.7381 | -102.5737
23.7418 | -102.5739
23.7455 | -102.5721
23.7492 | -102.5701
23.753 | -102.5691
23.7567 | -102.5681
23.7604 | -102.5685
23.7641 | -102.5715
23.7679 | -102.5723
23.7716 | -102.5733
23.7753 | -102.5765
23.7791 | -102.5808
23.7828 | -102.5857
23.7865 | -102.5913
23.7902 | -102.5978
23.794 | -102.6039
23.7977 | -102.6108
23.8014 | -102.6181
23.8051 | -102.6231
23.8089 | -102.6281
23.8126 | -102.6351
23.8163 | -102.6448
23.82 | -102.656
23.8238 | -102.6654
23.8275 | -102.6749
23.8312 | -102.684
23.8349 | -102.695
23.8387 | -102.7087
23.8424 | -102.7236
23.8461 | -102.7355
23.8498 | -102.7471
23.8536 | -102.7583
23.8573 | -102.7701
23.861 | -102.7814
23.8647 | -102.7919
23.8685 | -102.8018
23.8722 | -102.8132
23.8759 | -102.8242
23.8797 | -102.8322
23.8834 | -102.8413
23.8871 | -102.8504
23.8908 | -102.8582
23.8946 | -102.864
23.8983 | -102.87
23.902 | -102.8753
23.9057 | -102.8793
23.9095 | -102.8853
23.9132 | -102.8926
23.9169 | -102.8987
23.9206 | -102.9047
23.9244 | -102.9107
23.9281 | -102.9145
23.9318 | -102.9187
23.9355 | -102.9238
23.9393 | -102.928
23.943 | -102.9314
23.9467 | -102.9355
23.9504 | -102.9415
23.9542 | -102.9482
23.9579 | -102.9538
23.9616 | -102.9576
23.9653 | -102.9615
23.9691 | -102.9643
23.9728 | -102.9673
23.9765 | -102.9723
23.9803 | -102.9773
23.984 | -102.9833
23.9877 | -102.9892
23.9914 | -102.9953
23.9952 | -103.0003
23.9989 | -103.0048
24.0026 | -103.0081
24.0063 | -103.0124
24.0101 | -103.018
24.0138 | -103.0224
24.0175 | -103.0268
24.0212 | -103.0315
24.025 | -103.0375
24.0287 | -103.0474
24.0324 | -103.0561
24.0361 | -103.0636
24.0399 | -103.0739
24.0436 | -103.0836
24.0473 | -103.0934
24.051 | -103.1023
24.0548 | -103.1107
24.0585 | -103.1229
24.0622 | -103.1366
24.0659 | -103.1487
24.0697 | -103.158
24.0734 | -103.1657
24.0771 | -103.1742
24.0809 | -103.1849
24.0846 | -103.1955
24.0883 | -103.2052
24.092 | -103.2119
24.0958 | -103.219
24.0995 | -103.2267
24.1032 | -103.2335
24.1069 | -103.2442
24.1107 | -103.2554
24.1144 | -103.2621
24.1181 | -103.2697
24.1218 | -103.2779
24.1256 | -103.283
24.1293 | -103.2871
24.133 | -103.2925
24.1367 | -103.2971
24.1405 | -103.3031
24.1442 | -103.3118
24.1479 | -103.3188
24.1516 | -103.3222
24.1554 | -103.3222
24.1591 | -103.3214
24.1628 | -103.322
24.1666 | -103.322
24.1703 | -103.3242
24.174 | -103.3247
24.1777 | -103.3208
24.1815 | -103.3169
24.1852 | -103.3143
24.1889 | -103.3083
24.1926 | -103.3063
24.1964 | -103.3032
24.2001 | -103.2973
24.2038 | -103.2933
24.2075 | -103.2915
24.2113 | -103.2873
24.215 | -103.2825
24.2187 | -103.2793
24.2224 | -103.2744
24.2262 | -103.2672
24.2299 | -103.2629
24.2336 | -103.2583
24.2373 | -103.2526
24.2411 | -103.2453
24.2448 | -103.2379
24.2485 | -103.2282
24.2522 | -103.2109
24.256 | -103.1848
24.2597 | -103.1451
24.2634 | -103.0949
24.2672 | -103.0287
24.2709 | -102.965
24.2746 | -102.9034
24.2783 | -102.8236
24.2821 | -102.7443
24.2858 | -102.6832
24.2895 | -102.6273
24.2932 | -102.5688
24.297 | -102.5072
24.3007 | -102.4425
24.3044 | -102.3737
24.3081 | -102.3002
24.3119 | -102.2239
24.3156 | -102.1445
24.3193 | -102.0627
24.323 | -101.9791
24.3268 | -101.8942
24.3305 | -101.8069
24.3342 | -101.7171
24.3379 | -101.6257
24.3417 | -101.5323
24.3454 | -101.4402
24.3491 | -101.3472
24.3528 | -101.2523
24.3566 | -101.1583
24.3603 | -101.0668
24.364 | -100.9758
24.3678 | -100.8853
24.3715 | -100.7944
24.3752 | -100.7059
24.3789 | -100.6213
24.3827 | -100.5399
24.3864 | -100.4617
24.3901 | -100.3869
24.3938 | -100.3165
24.3976 | -100.2506
24.4013 | -100.1886
24.405 | -100.1315
24.4087 | -100.0791
24.4125 | -100.0323
24.4162 | -99.9913
24.4199 | -99.9548
24.4236 | -99.923
24.4274 | -99.8972
24.4311 | -99.8778
24.4348 | -99.8645
24.4385 | -99.8573
24.4423 | -99.8564
24.446 | -99.8605
24.4497 | -99.8707
24.4534 | -99.8873
24.4572 | -99.9099
24.4609 | -99.9375
24.4646 | -99.9703
24.4684 | -100.0097
24.4721 | -100.0546
24.4758 | -100.1039
24.4795 | -100.1582
24.4833 | -100.2177
24.487 | -100.281
24.4907 | -100.3483
24.4944 | -100.4203
24.4982 | -100.4966
24.5019 | -100.5766
24.5056 | -100.6592
24.5093 | -100.7452
24.5131 | -100.833
24.5168 | -100.9231
24.5205 | -101.0149
24.5242 | -101.1078
24.528 | -101.2022
24.5317 | -101.2954
24.5354 | -101.3881
24.5391 | -101.4814
24.5429 | -101.5745
24.5466 | -101.6654
24.5503 | -101.753
24.5541 | -101.8376
24.5578 | -101.9164
24.5615 | -101.9901
24.5652 | -102.0583
24.569 | -102.112
24.5727 | -102.1581
24.5764 | -102.1938
24.5801 | -102.2159
24.5839 | -102.2211
24.5876 | -102.2251
24.5913 | -102.2335
24.595 | -102.2325
24.5988 | -102.2297
24.6025 | -102.2284
24.6062 | -102.2284
24.6099 | -102.2253
24.6137 | -102.2192
24.6174 | -102.1961
24.6211 | -102.1592
24.6248 | -102.1219
24.6286 | -102.0692
24.6323 | -101.9922
24.636 | -101.9052
24.6397 | -101.8161
24.6435 | -101.7248
24.6472 | -101.6318
24.6509 | -101.5364
24.6547 | -101.4404
24.6584 | -101.3444
24.6621 | -101.2476
24.6658 | -101.1498
24.6696 | -101.0519
24.6733 | -100.9539
24.677 | -100.857
24.6807 | -100.7611
24.6845 | -100.6665
24.6882 | -100.5758
24.6919 | -100.4875
24.6956 | -100.402
24.6994 | -100.3217
24.7031 | -100.2451
24.7068 | -100.1711
24.7105 | -100.1002
24.7143 | -100.0333
24.718 | -99.9718
24.7217 | -99.9159
24.7254 | -99.9014
24.7292 | -99.8571
24.7329 | -99.817
24.7366 | -99.7823
24.7403 | -99.7528
24.7441 | -99.7292
24.7478 | -99.7106
24.7515 | -99.6968
24.7553 | -99.6887
24.759 | -99.6862
24.7627 | -99.6882
24.7664 | -99.695
24.7702 | -99.7086
24.7739 | -99.7282
24.7776 | -99.7521
24.7813 | -99.7809
24.7851 | -99.8152
24.7888 | -99.8544
24.7925 | -99.8978
24.7962 | -99.946
24.8 | -99.9985
24.8037 | -100.0547
24.8074 | -100.1149
24.8111 | -100.1797
24.8149 | -100.2483
24.8186 | -100.3196
24.8223 | -100.3932
24.826 | -100.4697
24.8298 | -100.5492
24.8335 | -100.6305
24.8372 | -100.7138
24.8409 | -100.7992
24.8447 | -100.8859
24.8484 | -100.9728
24.8521 | -101.0598
24.8559 | -101.1473
24.8596 | -101.2328
24.8633 | -101.3184
24.867 | -101.4038
24.8708 | -101.4894
24.8745 | -101.5746
24.8782 | -101.6577
24.8819 | -101.7405
24.8857 | -101.8224
24.8894 | -101.9012
24.8931 | -101.9776
24.8968 | -102.0527
24.9006 | -102.1266
24.9043 | -102.1991
24.908 | -102.2721
24.9117 | -102.344
24.9155 | -102.4153
24.9192 | -102.492
24.9229 | -102.5669
24.9266 | -102.6404
24.9304 | -102.7147
24.9341 | -102.7941
24.9378 | -102.8603
24.9415 | -102.9074
24.9453 | -102.9536
24.949 | -102.9945
24.9527 | -103.0288
24.9565 | -103.0594
24.9602 | -103.0869
24.9639 | -103.1148
24.9676 | -103.1418
24.9714 | -103.1656
24.9751 | -103.188
24.9788 | -103.2128
24.9825 | -103.2353
24.9863 | -103.2572
24.99 | -103.2758
24.9937 | -103.2916
24.9974 | -103.3053
25.0012 | -103.3204
25.0049 | -103.3355
25.0086 | -103.3474
25.0123 | -103.3581
25.0161 | -103.369
25.0198 | -103.3783
25.0235 | -103.3855
25.0272 | -103.3914
25.031 | -103.3989
25.0347 | -103.4063
25.0384 | -103.4113
25.0421 | -103.4144
25.0459 | -103.4143
25.0496 | -103.4099
25.0533 | -103.4076
25.0571 | -103.4055
25.0608 | -103.4041
25.0645 | -103.403
25.0682 | -103.4006
25.072 | -103.3963
25.0757 | -103.3933
25.0794 | -103.3879
25.0831 | -103.3813
25.0869 | -103.3758
25.0906 | -103.3704
25.0943 | -103.3645
25.098 | -103.3579
25.1018 | -103.3513
25.1055 | -103.3437
25.1092 | -103.3361
25.1129 | -103.3278
25.1167 | -103.3206
25.1204 | -103.3139
25.1241 | -103.3058
25.1278 | -103.2954
25.1316 | -103.287
25.1353 | -103.2795
25.139 | -103.2707
25.1428 | -103.2627
25.1465 | -103.2554
25.1502 | -103.2478
25.1539 | -103.2394
25.1577 | -103.2312
25.1614 | -103.2232
25.1651 | -103.2158
25.1688 | -103.2096
25.1726 | -103.2027
25.1763 | -103.1959
25.18 | -103.1914
25.1837 | -103.1861
25.1875 | -103.1803
25.1912 | -103.1749
25.1949 | -103.1693
25.1986 | -103.1635
25.2024 | -103.1585
25.2061 | -103.1539
25.2098 | -103.1486
25.2135 | -103.1459
25.2173 | -103.143
25.221 | -103.1382
25.2247 | -103.1323
25.2284 | -103.1274
25.2322 | -103.1243
25.2359 | -103.1194
25.2396 | -103.112
25.2434 | -103.1068
25.2471 | -103.1006
25.2508 | -103.0954
25.2545 | -103.0894
25.2583 | -103.0808
25.262 | -103.0665
25.2657 | -103.0583
25.2694 | -103.0526
25.2732 | -103.0448
25.2769 | -103.0389
25.2806 | -103.0324
25.2843 | -103.0245
25.2881 | -103.0187
25.2918 | -103.0125
25.2955 | -103.0071
25.2992 | -103.0023
25.303 | -102.9986
25.3067 | -102.995
25.3104 | -102.9908
25.3141 | -102.9869
25.3179 | -102.9826
25.3216 | -102.9774
25.3253 | -102.9737
25.329 | -102.9719
25.3328 | -102.9709
25.3365 | -102.9685
25.3402 | -102.9668
25.344 | -102.9648
25.3477 | -102.9629
25.3514 | -102.9624
25.3551 | -102.9626
25.3589 | -102.9618
25.3626 | -102.961
25.3663 | -102.9592
25.37 | -102.9568
25.3738 | -102.9557
25.3775 | -102.9551
25.3812 | -102.9572
25.3849 | -102.9599
25.3887 | -102.9629
25.3924 | -102.9653
25.3961 | -102.9674
25.3998 | -102.9693
25.4036 | -102.9708
25.4073 | -102.9752
25.411 | -102.9782
25.4147 | -102.9806
25.4185 | -102.9844
25.4222 | -102.9881
25.4259 | -102.9912
25.4296 | -102.9959
25.4334 | -102.9989
25.4371 | -103.0036
25.4408 | -103.0104
25.4446 | -103.0168
25.4483 | -103.0239
25.452 | -103.0309
25.4557 | -103.0372
25.4595 | -103.0433
25.4632 | -103.0513
25.4669 | -103.059
25.4706 | -103.0672
25.4744 | -103.0763
25.4781 | -103.0863
25.4818 | -103.0946
25.4855 | -103.1027
25.4893 | -103.11
25.493 | -103.1165
25.4967 | -103.1233
25.5004 | -103.1288
25.5042 | -103.135
25.5079 | -103.1411
25.5116 | -103.1471
25.5153 | -103.1534
25.5191 | -103.1591
25.5228 | -103.1651
25.5265 | -103.1701
25.5302 | -103.1776
25.534 | -103.1857
25.5377 | -103.1922
25.5414 | -103.1985
25.5452 | -103.2059
25.5489 | -103.2135
25.5526 | -103.2196
25.5563 | -103.226
25.5601 | -103.2339
25.5638 | -103.2414
25.5675 | -103.249
25.5712 | -103.2564
25.575 | -103.2655
25.5787 | -103.2762
25.5824 | -103.2875
25.5861 | -103.2975
25.5899 | -103.3062
25.5936 | -103.3165
25.5973 | -103.3289
25.601 | -103.3384
25.6048 | -103.3489
25.6085 | -103.3599
25.6122 | -103.3721
25.6159 | -103.3855
25.6197 | -103.4001
25.6234 | -103.4148
25.6271 | -103.4264
25.6309 | -103.4353
25.6346 | -103.445
25.6383 | -103.4535
25.642 | -103.4629
25.6458 | -103.4739
25.6495 | -103.4826
25.6532 | -103.4897
25.6569 | -103.4953
25.6607 | -103.5019
25.6644 | -103.5073
25.6681 | -103.5133
25.6718 | -103.5198
25.6756 | -103.5244
25.6793 | -103.5285
25.683 | -103.535
25.6867 | -103.5394
25.6905 | -103.5425
25.6942 | -103.5468
25.6979 | -103.5504
25.7016 | -103.5503
25.7054 | -103.5496
25.7091 | -103.5502
25.7128 | -103.5521
25.7165 | -103.5532
25.7203 | -103.5543
25.724 | -103.5546
25.7277 | -103.5566
25.7315 | -103.5592
25.7352 | -103.5592
25.7389 | -103.5589
25.7426 | -103.5571
25.7464 | -103.5559
25.7501 | -103.5532
25.7538 | -103.551
25.7575 | -103.5491
25.7613 | -103.5483
25.765 | -103.5502
25.7687 | -103.5509
25.7724 | -103.5512
25.7762 | -103.5531
25.7799 | -103.553
25.7836 | -103.5524
25.7873 | -103.5527
25.7911 | -103.5509
25.7948 | -103.5488
25.7985 | -103.5516
25.8022 | -103.5525
25.806 | -103.5533
25.8097 | -103.555
25.8134 | -103.5559
25.8171 | -103.5565
25.8209 | -103.5594
25.8246 | -103.5627
25.8283 | -103.5651
25.8321 | -103.5689
25.8358 | -103.5719
25.8395 | -103.5725
25.8432 | -103.5706
25.847 | -103.5686
25.8507 | -103.5709
25.8544 | -103.5748
25.8581 | -103.5788
25.8619 | -103.5824
25.8656 | -103.5847
25.8693 | -103.5877
25.873 | -103.592
25.8768 | -103.5951
25.8805 | -103.598
25.8842 | -103.6017
25.8879 | -103.6035
25.8917 | -103.6051
25.8954 | -103.6066
25.8991 | -103.609
25.9028 | -103.6104
25.9066 | -103.6119
25.9103 | -103.6142
25.914 | -103.6159
25.9178 | -103.6179
25.9215 | -103.6209
25.9252 | -103.6233
25.9289 | -103.6258
25.9327 | -103.6294
25.9364 | -103.6329
25.9401 | -103.6344
25.9438 | -103.6342
25.9476 | -103.6344
25.9513 | -103.6367
25.955 | -103.6379
25.9587 | -103.638
25.9625 | -103.6386
25.9662 | -103.6421
25.9699 | -103.6457
25.9736 | -103.6464
25.9774 | -103.6474
25.9811 | -103.6502
25.9848 | -103.6527
25.9885 | -103.6548
25.9923 | -103.6557
25.996 | -103.6573
25.9997 | -103.658
26.0034 | -103.6585
26.0072 | -103.6597
26.0109 | -103.662
26.0146 | -103.666
26.0184 | -103.6707
26.0221 | -103.6751
26.0258 | -103.6781
26.0295 | -103.6838
26.0333 | -103.6885
26.037 | -103.6934
26.0407 | -103.7008
26.0444 | -103.7084
26.0482 | -103.7142
26.0519 | -103.7208
26.0556 | -103.7275
26.0593 | -103.7327
26.0631 | -103.7397
26.0668 | -103.7487
26.0705 | -103.7565
26.0742 | -103.7636
26.078 | -103.77
26.0817 | -103.7777
26.0854 | -103.7851
26.0891 | -103.7931
26.0929 | -103.801
26.0966 | -103.8093
26.1003 | -103.8168
26.104 | -103.8239
26.1078 | -103.8306
26.1115 | -103.8376
26.1152 | -103.8446
26.119 | -103.8499
26.1227 | -103.8565
26.1264 | -103.8628
26.1301 | -103.8692
26.1339 | -103.8754
26.1376 | -103.8809
26.1413 | -103.8842
26.145 | -103.8915
26.1488 | -103.9008
26.1525 | -103.9082
26.1562 | -103.9161
26.1599 | -103.9257
26.1637 | -103.9338
26.1674 | -103.9408
26.1711 | -103.9477
26.1748 | -103.9541
26.1786 | -103.9584
26.1823 | -103.9642
26.186 | -103.9735
26.1897 | -103.9825
26.1935 | -103.9902
26.1972 | -103.9998
26.2009 | -104.0084
26.2046 | -104.0177
26.2084 | -104.0292
26.2121 | -104.0425
26.2158 | -104.0562
26.2196 | -104.0654
26.2233 | -104.0748
26.227 | -104.0826
26.2307 | -104.0907
26.2345 | -104.1013
26.2382 | -104.1101
26.2419 | -104.1181
26.2456 | -104.1279
26.2494 | -104.1358
26.2531 | -104.1425
26.2568 | -104.1492
26.2605 | -104.1532
26.2643 | -104.1577
26.268 | -104.1606
26.2717 | -104.164
26.2754 | -104.1681
26.2792 | -104.1703
26.2829 | -104.1715
26.2866 | -104.1724
26.2903 | -104.1717
26.2941 | -104.1716
26.2978 | -104.1721
26.3015 | -104.1719
26.3052 | -104.171
26.309 | -104.1703
26.3127 | -104.1712
26.3164 | -104.1696
26.3202 | -104.1645
26.3239 | -104.1623
26.3276 | -104.1611
26.3313 | -104.1583
26.3351 | -104.1537
26.3388 | -104.1476
26.3425 | -104.1439
26.3462 | -104.1415
26.35 | -104.1374
26.3537 | -104.1319
26.3574 | -104.1266
26.3611 | -104.1231
26.3649 | -104.1219
26.3686 | -104.1196
26.3723 | -104.1186
26.376 | -104.1187
26.3798 | -104.1174
26.3835 | -104.1149
26.3872 | -104.1112
26.3909 | -104.1092
26.3947 | -104.1072
26.3984 | -104.1058
26.4021 | -104.1036
26.4059 | -104.1029
26.4096 | -104.1017
26.4133 | -104.0999
26.417 | -104.0989
26.4208 | -104.098
26.4245 | -104.0965
26.4282 | -104.0936
26.4319 | -104.0933
26.4357 | -104.094
26.4394 | -104.0972
26.4431 | -104.1011
26.4468 | -104.1039
26.4506 | -104.1068
26.4543 | -104.1101
26.458 | -104.1146
26.4617 | -104.1181
26.4655 | -104.1213
26.4692 | -104.1241
26.4729 | -104.1259
26.4766 | -104.1294
26.4804 | -104.1318
26.4841 | -104.1348
26.4878 | -104.1376
26.4915 | -104.1417
26.4953 | -104.1428
26.499 | -104.144
26.5027 | -104.146
26.5065 | -104.1485
26.5102 | -104.1513
26.5139 | -104.1547
26.5176 | -104.1535
26.5214 | -104.1512
26.5251 | -104.1505
26.5288 | -104.1496
26.5325 | -104.1467
26.5363 | -104.1446
26.54 | -104.1428
26.5437 | -104.1403
26.5474 | -104.1356
26.5512 | -104.1328
26.5549 | -104.1291
26.5586 | -104.1227
26.5623 | -104.1157
26.5661 | -104.1095
26.5698 | -104.1042
26.5735 | -104.0983
26.5772 | -104.0899
26.581 | -104.0804
26.5847 | -104.0726
26.5884 | -104.065
26.5921 | -104.0584
26.5959 | -104.0505
26.5996 | -104.0426
26.6033 | -104.0329
26.6071 | -104.0247
26.6108 | -104.0174
26.6145 | -104.0104
26.6182 | -104.0029
26.622 | -103.9972
26.6257 | -103.9921
26.6294 | -103.9859
26.6331 | -103.9807
26.6369 | -103.9742
26.6406 | -103.9679
26.6443 | -103.963
26.648 | -103.9591
26.6518 | -103.9547
26.6555 | -103.9503
26.6592 | -103.9478
26.6629 | -103.9467
26.6667 | -103.944
26.6704 | -103.9408
26.6741 | -103.938
26.6778 | -103.9362
26.6816 | -103.9354
26.6853 | -103.9352
26.689 | -103.936
26.6927 | -103.9379
26.6965 | -103.9381
26.7002 | -103.9383
26.7039 | -103.9382
26.7077 | -103.9385
26.7114 | -103.9389
26.7151 | -103.9369
26.7188 | -103.9337
26.7226 | -103.9306
26.7263 | -103.9292
26.73 | -103.9292
26.7337 | -103.9297
26.7375 | -103.9301
26.7412 | -103.9306
26.7449 | -103.9291
26.7486 | -103.9277
26.7524 | -103.9272
26.7561 | -103.9276
26.7598 | -103.9281
26.7635 | -103.93
26.7673 | -103.9302
26.771 | -103.9296
26.7747 | -103.9303
26.7784 | -103.9307
26.7822 | -103.9303
26.7859 | -103.9311
26.7896 | -103.9336
26.7933 | -103.9348
26.7971 | -103.9358
26.8008 | -103.9381
26.8045 | -103.9401
26.8083 | -103.94
26.812 | -103.9374
26.8157 | -103.9339
26.8194 | -103.9317
26.8232 | -103.9314
26.8269 | -103.931
26.8306 | -103.9302
26.8343 | -103.9284
26.8381 | -103.9264
26.8418 | -103.925
26.8455 | -103.9252
26.8492 | -103.9261
26.853 | -103.9249
26.8567 | -103.9251
26.8604 | -103.9243
26.8641 | -103.9216
26.8679 | -103.9161
26.8716 | -103.9102
26.8753 | -103.9011
26.879 | -103.8981
26.8828 | -103.8929
26.8865 | -103.8868
26.8902 | -103.8798
26.894 | -103.8724
26.8977 | -103.8673
26.9014 | -103.8634
26.9051 | -103.8576
26.9089 | -103.8509
26.9126 | -103.845
26.9163 | -103.8377
26.92 | -103.8297
26.9238 | -103.8219
26.9275 | -103.8164
26.9312 | -103.8085
26.9349 | -103.7987
26.9387 | -103.7893
26.9424 | -103.7795
26.9461 | -103.771
26.9498 | -103.7596
26.9536 | -103.7473
26.9573 | -103.7368
26.961 | -103.7272
26.9647 | -103.7165
26.9685 | -103.7045
26.9722 | -103.692
26.9759 | -103.6792
26.9796 | -103.6664
26.9834 | -103.6555
26.9871 | -103.6435
26.9908 | -103.6313
26.9946 | -103.6179
26.9983 | -103.6034
27.002 | -103.5877
27.0057 | -103.5745
27.0095 | -103.5611
27.0132 | -103.5481
27.0169 | -103.5351
27.0206 | -103.5233
27.0244 | -103.5096
27.0281 | -103.4945
27.0318 | -103.4818
27.0355 | -103.4709
27.0393 | -103.4585
27.043 | -103.4457
27.0467 | -103.4306
27.0504 | -103.4144
27.0542 | -103.4005
27.0579 | -103.388
27.0616 | -103.3758
27.0653 | -103.3626
27.0691 | -103.3498
27.0728 | -103.3382
27.0765 | -103.3268
27.0802 | -103.3159
27.084 | -103.3058
27.0877 | -103.2963
27.0914 | -103.2868
27.0952 | -103.2793
27.0989 | -103.2732
27.1026 | -103.2676
27.1063 | -103.262
27.1101 | -103.2545
27.1138 | -103.2466
27.1175 | -103.2392
27.1212 | -103.2317
27.125 | -103.2258
27.1287 | -103.2207
27.1324 | -103.2161
27.1361 | -103.2106
27.1399 | -103.205
27.1436 | -103.1983
27.1473 | -103.1947
27.151 | -103.1923
27.1548 | -103.189
27.1585 | -103.1861
27.1622 | -103.1806
27.1659 | -103.1733
27.1697 | -103.1677
27.1734 | -103.1628
27.1771 | -103.1579
27.1808 | -103.1516
27.1846 | -103.1473
27.1883 | -103.1451
27.192 | -103.1438
27.1958 | -103.1409
27.1995 | -103.1368
27.2032 | -103.1331
27.2069 | -103.1282
27.2107 | -103.124
27.2144 | -103.1208
27.2181 | -103.1189
27.2218 | -103.1159
27.2256 | -103.1119
27.2293 | -103.1061
27.233 | -103.1012
27.2367 | -103.0975
27.2405 | -103.0914
27.2442 | -103.0878
27.2479 | -103.0858
27.2516 | -103.0824
27.2554 | -103.076
27.2591 | -103.068
27.2628 | -103.0591
27.2665 | -103.0517
27.2703 | -103.0427
27.274 | -103.0339
27.2777 | -103.0251
27.2814 | -103.0159
27.2852 | -103.0055
27.2889 | -102.9949
27.2926 | -102.9846
27.2964 | -102.9738
27.3001 | -102.963
27.3038 | -102.9514
27.3075 | -102.9376
27.3113 | -102.9225
27.315 | -102.9103
27.3187 | -102.9002
27.3224 | -102.8886
27.3262 | -102.8754
27.3299 | -102.863
27.3336 | -102.8504
27.3373 | -102.8384
27.3411 | -102.8267
27.3448 | -102.8154
27.3485 | -102.803
27.3522 | -102.7909
27.356 | -102.7776
27.3597 | -102.7632
27.3634 | -102.7488
27.3671 | -102.7344
27.3709 | -102.7197
27.3746 | -102.7046
27.3783 | -102.6916
27.382 | -102.6834
27.3858 | -102.6756
27.3895 | -102.6652
27.3932 | -102.6554
27.397 | -102.6454
27.4007 | -102.6368
27.4044 | -102.6276
27.4081 | -102.6196
27.4119 | -102.6195
27.4156 | -102.6121
27.4193 | -102.6045
27.423 | -102.5974
27.4268 | -102.5923
27.4305 | -102.5882
27.4342 | -102.5865
27.4379 | -102.5864
27.4417 | -102.5854
27.4454 | -102.5827
27.4491 | -102.5736
27.4528 | -102.5605
27.4566 | -102.5492
27.4603 | -102.5424
27.464 | -102.5391
27.4677 | -102.5359
27.4715 | -102.5309
27.4752 | -102.5263
27.4789 | -102.5227
27.4827 | -102.52
27.4864 | -102.5176
27.4901 | -102.5157
27.4938 | -102.5134
27.4976 | -102.5125
27.5013 | -102.511
27.505 | -102.5089
27.5087 | -102.5066
27.5125 | -102.5036
27.5162 | -102.4994
27.5199 | -102.4955
27.5236 | -102.4931
27.5274 | -102.4897
27.5311 | -102.4872
27.5348 | -102.4827
27.5385 | -102.4778
27.5423 | -102.4718
27.546 | -102.4667
27.5497 | -102.4639
27.5534 | -102.4612
27.5572 | -102.455
27.5609 | -102.4484
27.5646 | -102.4428
27.5683 | -102.4378
27.5721 | -102.4317
27.5758 | -102.4232
27.5795 | -102.4152
27.5833 | -102.406
27.587 | -102.3987
27.5907 | -102.391
27.5944 | -102.3835
27.5982 | -102.3756
27.6019 | -102.3681
27.6056 | -102.362
27.6093 | -102.3572
27.6131 | -102.352
27.6168 | -102.3484
27.6205 | -102.3453
27.6242 | -102.3408
27.628 | -102.3392
27.6317 | -102.3378
27.6354 | -102.3365
27.6391 | -102.3374
27.6429 | -102.337
27.6466 | -102.3328
27.6503 | -102.3322
27.654 | -102.3325
27.6578 | -102.3351
27.6615 | -102.3385
27.6652 | -102.343
27.6689 | -102.3468
27.6727 | -102.3509
27.6764 | -102.3542
27.6801 | -102.3566
27.6839 | -102.3588
27.6876 | -102.3634
27.6913 | -102.3677
27.695 | -102.3743
27.6988 | -102.3808
27.7025 | -102.3876
27.7062 | -102.3927
27.7099 | -102.3986
27.7137 | -102.405
27.7174 | -102.41
27.7211 | -102.4149
27.7248 | -102.42
27.7286 | -102.4246
27.7323 | -102.4281
27.736 | -102.4325
27.7397 | -102.4368
27.7435 | -102.4413
27.7472 | -102.4461
27.7509 | -102.4491
27.7546 | -102.4519
27.7584 | -102.4525
27.7621 | -102.4508
27.7658 | -102.4481
27.7695 | -102.4437
27.7733 | -102.439
27.777 | -102.4368
27.7807 | -102.435
27.7845 | -102.4321
27.7882 | -102.4287
27.7919 | -102.4267
27.7956 | -102.4228
27.7994 | -102.418
27.8031 | -102.4135
27.8068 | -102.4079
27.8105 | -102.4027
27.8143 | -102.3981
27.818 | -102.393
27.8217 | -102.3873
27.8254 | -102.3811
27.8292 | -102.3748
27.8329 | -102.3699
27.8366 | -102.3655
27.8403 | -102.3672
27.8441 | -102.3721
27.8478 | -102.3752
27.8515 | -102.3711
27.8552 | -102.3636
27.859 | -102.3544
27.8627 | -102.3487
27.8664 | -102.3435
27.8701 | -102.3378
27.8739 | -102.3329
27.8776 | -102.3266
27.8813 | -102.3197
27.8851 | -102.3119
27.8888 | -102.3029
27.8925 | -102.2944
27.8962 | -102.2876
27.9 | -102.2804
27.9037 | -102.2761
27.9074 | -102.2705
27.9111 | -102.2636
27.9149 | -102.2562
27.9186 | -102.2509
27.9223 | -102.2436
27.926 | -102.2347
27.9298 | -102.224
27.9335 | -102.2158
27.9372 | -102.2079
27.9409 | -102.1983
27.9447 | -102.1895
27.9484 | -102.1793
27.9521 | -102.1722
27.9558 | -102.1648
27.9596 | -102.1555
27.9633 | -102.146
27.967 | -102.1378
27.9708 | -102.1302
27.9745 | -102.1206
27.9782 | -102.1124
27.9819 | -102.1011
27.9857 | -102.0902
27.9894 | -102.0812
27.9931 | -102.071
27.9968 | -102.0586
28.0006 | -102.0455