## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 281 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 264
Hood | 02-92B/P | 191
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1003.4 | hPa |  | PRC PRO 01
Humidity | 54.1 | %/Hr
Internal temperature of calibration bench | 24.14 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1013 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 25.1 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.09 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.99 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3126 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 17/09/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 19/09/2018
Technical Manager | Anthony HUE | Approval |  | 27/09/2018
CEO | Robert GLEMOT | Approval |  | 27/09/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4165 | 57.5332
0.012209 | -35.9851 | 38.4123
0.018313 | -34.9146 | 28.2307
0.024417 | -34.4481 | 21.9893
0.030521 | -34.335 | 18.2144
0.036626 | -34.1915 | 15.1548
0.04273 | -34.0964 | 12.773
0.048834 | -34.0495 | 11.2458
0.054938 | -33.9861 | 10.1138
0.061043 | -33.9281 | 9.0054
0.067147 | -33.9175 | 8.1767
0.073251 | -33.9197 | 7.4187
0.079355 | -33.9393 | 6.7667
0.08546 | -33.9534 | 6.183
0.091564 | -33.9482 | 5.7419
0.097668 | -33.9309 | 5.477
0.10377 | -33.9114 | 5.2187
0.10988 | -33.8955 | 4.941
0.11598 | -33.8866 | 4.6586
0.12209 | -33.8875 | 4.389
0.12819 | -33.8963 | 4.116
0.13429 | -33.9066 | 3.8322
0.1404 | -33.9097 | 3.5499
0.1465 | -33.9112 | 3.2888
0.15261 | -33.9081 | 3.0623
0.15871 | -33.9015 | 2.8919
0.16482 | -33.8939 | 2.7794
0.17092 | -33.8927 | 2.7311
0.17702 | -33.8914 | 2.6518
0.18313 | -33.8912 | 2.5687
0.18923 | -33.8912 | 2.4583
0.19534 | -33.8901 | 2.3331
0.20144 | -33.8886 | 2.1942
0.20754 | -33.8856 | 2.0759
0.21365 | -33.8845 | 1.9631
0.21975 | -33.8839 | 1.8653
0.22586 | -33.885 | 1.7778
0.23196 | -33.8875 | 1.7041
0.23807 | -33.8918 | 1.6399
0.24417 | -33.8962 | 1.5798
0.25027 | -33.9004 | 1.5161
0.25638 | -33.9031 | 1.4534
0.26248 | -33.9053 | 1.3872
0.26859 | -33.9065 | 1.3228
0.27469 | -33.9071 | 1.2535
0.2808 | -33.9081 | 1.1824
0.2869 | -33.909 | 1.1138
0.293 | -33.9099 | 1.0491
0.29911 | -33.9108 | 0.99136
0.30521 | -33.9116 | 0.93439
0.31132 | -33.9123 | 0.89735
0.31742 | -33.9134 | 0.85616
0.32353 | -33.9137 | 0.82204
0.32963 | -33.9146 | 0.79328
0.33573 | -33.9157 | 0.77164
0.34184 | -33.9171 | 0.74834
0.34794 | -33.9193 | 0.7213
0.35405 | -33.9215 | 0.69448
0.36015 | -33.9241 | 0.6661
0.36626 | -33.9262 | 0.63552
0.37236 | -33.9284 | 0.59627
0.37846 | -33.9305 | 0.54951
0.38457 | -33.9314 | 0.49194
0.39067 | -33.9324 | 0.4441
0.39678 | -33.9331 | 0.39183
0.40288 | -33.9335 | 0.33811
0.40899 | -33.9333 | 0.28704
0.41509 | -33.9326 | 0.23756
0.42119 | -33.9325 | 0.19638
0.4273 | -33.9324 | 0.15182
0.4334 | -33.9323 | 0.10741
0.43951 | -33.9325 | 0.06552
0.44561 | -33.9333 | 0.027098
0.45172 | -33.9349 | -0.0045592
0.45782 | -33.9357 | -0.040999
0.46392 | -33.9369 | -0.076826
0.47003 | -33.938 | -0.10921
0.47613 | -33.9389 | -0.14334
0.48224 | -33.9395 | -0.17841
0.48834 | -33.9388 | -0.21359
0.49445 | -33.9388 | -0.24905
0.50055 | -33.9384 | -0.28136
0.50665 | -33.9379 | -0.31599
0.51276 | -33.9371 | -0.35041
0.51886 | -33.9363 | -0.38448
0.52497 | -33.9353 | -0.40913
0.53107 | -33.9347 | -0.43893
0.53717 | -33.9342 | -0.46737
0.54328 | -33.9336 | -0.49088
0.54938 | -33.9335 | -0.51389
0.55549 | -33.9337 | -0.53541
0.56159 | -33.9348 | -0.54801
0.5677 | -33.9352 | -0.56353
0.5738 | -33.9363 | -0.58078
0.5799 | -33.9371 | -0.59659
0.58601 | -33.9382 | -0.60796
0.59211 | -33.9398 | -0.61568
0.59822 | -33.9412 | -0.62645
0.60432 | -33.9427 | -0.63961
0.61043 | -33.9439 | -0.64883
0.61653 | -33.9451 | -0.66014
0.62263 | -33.9466 | -0.67057
0.62874 | -33.9472 | -0.68195
0.63484 | -33.9481 | -0.69215
0.64095 | -33.9486 | -0.70318
0.64705 | -33.949 | -0.71382
0.65316 | -33.9492 | -0.72693
0.65926 | -33.9493 | -0.74334
0.66536 | -33.9483 | -0.76116
0.67147 | -33.948 | -0.77498
0.67757 | -33.9477 | -0.7898
0.68368 | -33.947 | -0.809
0.68978 | -33.9463 | -0.8265
0.69589 | -33.9455 | -0.84574
0.70199 | -33.9447 | -0.86858
0.70809 | -33.9439 | -0.88793
0.7142 | -33.9433 | -0.90865
0.7203 | -33.9423 | -0.92796
0.72641 | -33.9418 | -0.95052
0.73251 | -33.9412 | -0.97063
0.73862 | -33.9415 | -0.99227
0.74472 | -33.9413 | -1.0124
0.75082 | -33.9411 | -1.0345
0.75693 | -33.9415 | -1.0563
0.76303 | -33.9421 | -1.0782
0.76914 | -33.9439 | -1.0986
0.77524 | -33.9449 | -1.1183
0.78135 | -33.9462 | -1.1384
0.78745 | -33.9476 | -1.1581
0.79355 | -33.9489 | -1.1779
0.79966 | -33.9506 | -1.1973
0.80576 | -33.9528 | -1.2169
0.81187 | -33.9544 | -1.2384
0.81797 | -33.9563 | -1.2595
0.82408 | -33.958 | -1.2799
0.83018 | -33.9597 | -1.3004
0.83628 | -33.9615 | -1.319
0.84239 | -33.9627 | -1.3394
0.84849 | -33.9641 | -1.3609
0.8546 | -33.9655 | -1.3807
0.8607 | -33.9667 | -1.4013
0.86681 | -33.9678 | -1.4202
0.87291 | -33.9687 | -1.4391
0.87901 | -33.9692 | -1.4572
0.88512 | -33.9699 | -1.4753
0.89122 | -33.9706 | -1.4936
0.89733 | -33.9711 | -1.5094
0.90343 | -33.9715 | -1.5271
0.90953 | -33.9717 | -1.5416
0.91564 | -33.9718 | -1.5576
0.92174 | -33.972 | -1.5727
0.92785 | -33.9722 | -1.5874
0.93395 | -33.9723 | -1.6014
0.94006 | -33.9725 | -1.6145
0.94616 | -33.9724 | -1.6266
0.95226 | -33.9725 | -1.6395
0.95837 | -33.9727 | -1.6535
0.96447 | -33.9729 | -1.6665
0.97058 | -33.9729 | -1.6801
0.97668 | -33.9727 | -1.6936
0.98279 | -33.9726 | -1.7064
0.98889 | -33.9725 | -1.7173
0.99499 | -33.9725 | -1.7308
1.0011 | -33.9725 | -1.7418
1.0072 | -33.9722 | -1.7536
1.0133 | -33.9719 | -1.7646
1.0194 | -33.9716 | -1.7777
1.0255 | -33.9714 | -1.7915
1.0316 | -33.9709 | -1.8036
1.0377 | -33.9706 | -1.8174
1.0438 | -33.9701 | -1.8302
1.0499 | -33.9694 | -1.8456
1.056 | -33.9688 | -1.86
1.0621 | -33.9683 | -1.8745
1.0682 | -33.9679 | -1.8872
1.0743 | -33.9674 | -1.9027
1.0805 | -33.967 | -1.9173
1.0866 | -33.9661 | -1.9326
1.0927 | -33.9654 | -1.9473
1.0988 | -33.9649 | -1.9624
1.1049 | -33.9642 | -1.9769
1.111 | -33.9635 | -1.9925
1.1171 | -33.9629 | -2.0072
1.1232 | -33.9625 | -2.0228
1.1293 | -33.9622 | -2.0372
1.1354 | -33.962 | -2.0526
1.1415 | -33.9617 | -2.0684
1.1476 | -33.9611 | -2.0843
1.1537 | -33.9607 | -2.0993
1.1598 | -33.9608 | -2.1146
1.1659 | -33.9607 | -2.1284
1.172 | -33.9605 | -2.1427
1.1781 | -33.9603 | -2.1571
1.1842 | -33.9602 | -2.1717
1.1903 | -33.9606 | -2.1832
1.1964 | -33.9605 | -2.197
1.2025 | -33.9607 | -2.2114
1.2086 | -33.961 | -2.2243
1.2147 | -33.9611 | -2.2373
1.2209 | -33.9615 | -2.249
1.227 | -33.9622 | -2.2602
1.2331 | -33.9626 | -2.2726
1.2392 | -33.963 | -2.2849
1.2453 | -33.9634 | -2.2964
1.2514 | -33.9639 | -2.3068
1.2575 | -33.9644 | -2.3181
1.2636 | -33.9653 | -2.3273
1.2697 | -33.9659 | -2.337
1.2758 | -33.9665 | -2.3485
1.2819 | -33.9673 | -2.3588
1.288 | -33.9679 | -2.3698
1.2941 | -33.9687 | -2.3787
1.3002 | -33.9694 | -2.387
1.3063 | -33.97 | -2.3964
1.3124 | -33.9706 | -2.4057
1.3185 | -33.9713 | -2.4151
1.3246 | -33.972 | -2.4245
1.3307 | -33.9728 | -2.4333
1.3368 | -33.9736 | -2.4422
1.3429 | -33.9742 | -2.4514
1.349 | -33.9747 | -2.461
1.3551 | -33.9754 | -2.4706
1.3613 | -33.976 | -2.4791
1.3674 | -33.9766 | -2.4893
1.3735 | -33.9771 | -2.4992
1.3796 | -33.9775 | -2.5086
1.3857 | -33.9779 | -2.518
1.3918 | -33.9784 | -2.5285
1.3979 | -33.979 | -2.5379
1.404 | -33.9794 | -2.5495
1.4101 | -33.9799 | -2.5601
1.4162 | -33.9803 | -2.5709
1.4223 | -33.9808 | -2.5825
1.4284 | -33.9812 | -2.5928
1.4345 | -33.9815 | -2.6045
1.4406 | -33.9818 | -2.6182
1.4467 | -33.9822 | -2.6299
1.4528 | -33.9827 | -2.6415
1.4589 | -33.9832 | -2.6535
1.465 | -33.9834 | -2.6647
1.4711 | -33.9839 | -2.6763
1.4772 | -33.9844 | -2.6898
1.4833 | -33.9849 | -2.703
1.4894 | -33.9855 | -2.7154
1.4955 | -33.986 | -2.7276
1.5016 | -33.9865 | -2.7403
1.5078 | -33.9869 | -2.7543
1.5139 | -33.9874 | -2.7676
1.52 | -33.9878 | -2.7804
1.5261 | -33.9881 | -2.7932
1.5322 | -33.9885 | -2.8073
1.5383 | -33.9889 | -2.8201
1.5444 | -33.9894 | -2.833
1.5505 | -33.9899 | -2.8466
1.5566 | -33.9904 | -2.8598
1.5627 | -33.9909 | -2.8737
1.5688 | -33.9915 | -2.8875
1.5749 | -33.9919 | -2.9004
1.581 | -33.9922 | -2.914
1.5871 | -33.9926 | -2.9278
1.5932 | -33.993 | -2.9409
1.5993 | -33.9934 | -2.9553
1.6054 | -33.9938 | -2.9688
1.6115 | -33.9943 | -2.983
1.6176 | -33.9946 | -2.9959
1.6237 | -33.9948 | -3.0097
1.6298 | -33.9952 | -3.0225
1.6359 | -33.9956 | -3.0355
1.642 | -33.996 | -3.0478
1.6482 | -33.9962 | -3.0611
1.6543 | -33.9966 | -3.0749
1.6604 | -33.9969 | -3.0872
1.6665 | -33.9971 | -3.1
1.6726 | -33.9973 | -3.1128
1.6787 | -33.9977 | -3.1257
1.6848 | -33.9977 | -3.1386
1.6909 | -33.998 | -3.1516
1.697 | -33.9982 | -3.1641
1.7031 | -33.9983 | -3.177
1.7092 | -33.9983 | -3.1895
1.7153 | -33.9986 | -3.2018
1.7214 | -33.9984 | -3.2143
1.7275 | -33.9984 | -3.2272
1.7336 | -33.9984 | -3.2404
1.7397 | -33.9986 | -3.253
1.7458 | -33.9986 | -3.2655
1.7519 | -33.9987 | -3.2779
1.758 | -33.9988 | -3.2902
1.7641 | -33.999 | -3.3021
1.7702 | -33.999 | -3.3153
1.7763 | -33.999 | -3.3272
1.7824 | -33.999 | -3.3398
1.7885 | -33.999 | -3.3525
1.7947 | -33.9992 | -3.3648
1.8008 | -33.9992 | -3.3764
1.8069 | -33.9993 | -3.3884
1.813 | -33.9994 | -3.4008
1.8191 | -33.9994 | -3.4137
1.8252 | -33.9996 | -3.4267
1.8313 | -33.9997 | -3.4391
1.8374 | -33.9999 | -3.4512
1.8435 | -34 | -3.4634
1.8496 | -34.0001 | -3.4766
1.8557 | -34.0003 | -3.4895
1.8618 | -34.0007 | -3.5019
1.8679 | -34.0008 | -3.5141
1.874 | -34.001 | -3.5263
1.8801 | -34.0012 | -3.5388
1.8862 | -34.0014 | -3.5509
1.8923 | -34.0016 | -3.5633
1.8984 | -34.0018 | -3.5754
1.9045 | -34.0019 | -3.5872
1.9106 | -34.002 | -3.5995
1.9167 | -34.0021 | -3.6118
1.9228 | -34.0023 | -3.6238
1.9289 | -34.0024 | -3.6358
1.9351 | -34.0026 | -3.6468
1.9412 | -34.0029 | -3.6582
1.9473 | -34.0032 | -3.6704
1.9534 | -34.0034 | -3.682
1.9595 | -34.0036 | -3.6939
1.9656 | -34.0039 | -3.7062
1.9717 | -34.0042 | -3.7181
1.9778 | -34.0045 | -3.7296
1.9839 | -34.0048 | -3.7412
1.99 | -34.005 | -3.7523
1.9961 | -34.0051 | -3.7646
2.0022 | -34.0053 | -3.7761
2.0083 | -34.0055 | -3.7874
2.0144 | -34.0057 | -3.799
2.0205 | -34.006 | -3.8108
2.0266 | -34.0064 | -3.8223
2.0327 | -34.0066 | -3.834
2.0388 | -34.0071 | -3.8456
2.0449 | -34.0073 | -3.8571
2.051 | -34.0075 | -3.8686
2.0571 | -34.0079 | -3.8801
2.0632 | -34.0081 | -3.8917
2.0693 | -34.0083 | -3.904
2.0754 | -34.0086 | -3.9155
2.0816 | -34.0089 | -3.9277
2.0877 | -34.0091 | -3.9394
2.0938 | -34.0094 | -3.9516
2.0999 | -34.0096 | -3.9634
2.106 | -34.0098 | -3.9752
2.1121 | -34.01 | -3.9873
2.1182 | -34.0103 | -3.9982
2.1243 | -34.0105 | -4.0092
2.1304 | -34.0107 | -4.0205
2.1365 | -34.011 | -4.0323
2.1426 | -34.0112 | -4.0444
2.1487 | -34.0115 | -4.0558
2.1548 | -34.0119 | -4.0678
2.1609 | -34.0123 | -4.0794
2.167 | -34.0127 | -4.0908
2.1731 | -34.013 | -4.1029
2.1792 | -34.0132 | -4.1141
2.1853 | -34.0136 | -4.1262
2.1914 | -34.0139 | -4.1377
2.1975 | -34.0143 | -4.1493
2.2036 | -34.0145 | -4.1615
2.2097 | -34.0147 | -4.1731
2.2158 | -34.0151 | -4.1839
2.222 | -34.0153 | -4.1946
2.2281 | -34.0156 | -4.2061
2.2342 | -34.0159 | -4.2176
2.2403 | -34.0161 | -4.2285
2.2464 | -34.0163 | -4.2394
2.2525 | -34.0166 | -4.251
2.2586 | -34.0169 | -4.2618
2.2647 | -34.0172 | -4.2725
2.2708 | -34.0175 | -4.2839
2.2769 | -34.0178 | -4.2947
2.283 | -34.018 | -4.3061
2.2891 | -34.0182 | -4.3171
2.2952 | -34.0184 | -4.3279
2.3013 | -34.0186 | -4.339
2.3074 | -34.0189 | -4.3496
2.3135 | -34.0191 | -4.3612
2.3196 | -34.0193 | -4.3716
2.3257 | -34.0195 | -4.3824
2.3318 | -34.0197 | -4.3936
2.3379 | -34.0199 | -4.4046
2.344 | -34.02 | -4.4162
2.3501 | -34.0202 | -4.4271
2.3562 | -34.0204 | -4.4377
2.3623 | -34.0205 | -4.4483
2.3685 | -34.0206 | -4.459
2.3746 | -34.0208 | -4.4704
2.3807 | -34.021 | -4.4811
2.3868 | -34.0212 | -4.492
2.3929 | -34.0214 | -4.5029
2.399 | -34.0215 | -4.5135
2.4051 | -34.0218 | -4.5242
2.4112 | -34.0219 | -4.5351
2.4173 | -34.0221 | -4.5459
2.4234 | -34.0222 | -4.5568
2.4295 | -34.0224 | -4.5679
2.4356 | -34.0226 | -4.5789
2.4417 | -34.0227 | -4.5902
2.4478 | -34.0228 | -4.6013
2.4539 | -34.023 | -4.6118
2.46 | -34.0231 | -4.6234
2.4661 | -34.0233 | -4.6347
2.4722 | -34.0235 | -4.6451
2.4783 | -34.0236 | -4.6557
2.4844 | -34.0237 | -4.667
2.4905 | -34.0238 | -4.6781
2.4966 | -34.024 | -4.689
2.5027 | -34.0241 | -4.6997
2.5089 | -34.0243 | -4.7108
2.515 | -34.0243 | -4.7218
2.5211 | -34.0244 | -4.7328
2.5272 | -34.0246 | -4.7439
2.5333 | -34.0247 | -4.755
2.5394 | -34.0249 | -4.766
2.5455 | -34.025 | -4.7775
2.5516 | -34.0252 | -4.7881
2.5577 | -34.0254 | -4.7997
2.5638 | -34.0256 | -4.8117
2.5699 | -34.0257 | -4.8233
2.576 | -34.0258 | -4.8351
2.5821 | -34.0259 | -4.8465
2.5882 | -34.026 | -4.8578
2.5943 | -34.0262 | -4.8689
2.6004 | -34.0264 | -4.8801
2.6065 | -34.0265 | -4.8907
2.6126 | -34.0267 | -4.9014
2.6187 | -34.0269 | -4.9122
2.6248 | -34.027 | -4.9232
2.6309 | -34.0272 | -4.9349
2.637 | -34.0275 | -4.9467
2.6431 | -34.0276 | -4.9575
2.6492 | -34.0278 | -4.969
2.6554 | -34.028 | -4.9801
2.6615 | -34.0281 | -4.9913
2.6676 | -34.0284 | -5.003
2.6737 | -34.0287 | -5.0144
2.6798 | -34.029 | -5.0257
2.6859 | -34.0292 | -5.0373
2.692 | -34.0294 | -5.0484
2.6981 | -34.0296 | -5.0597
2.7042 | -34.0298 | -5.0711
2.7103 | -34.03 | -5.082
2.7164 | -34.0304 | -5.0932
2.7225 | -34.0306 | -5.1048
2.7286 | -34.0309 | -5.1161
2.7347 | -34.0311 | -5.1269
2.7408 | -34.0314 | -5.1372
2.7469 | -34.0317 | -5.1487
2.753 | -34.032 | -5.1596
2.7591 | -34.0323 | -5.1711
2.7652 | -34.0325 | -5.1828
2.7713 | -34.0327 | -5.1939
2.7774 | -34.033 | -5.2051
2.7835 | -34.0332 | -5.2163
2.7896 | -34.0334 | -5.2271
2.7958 | -34.0337 | -5.2381
2.8019 | -34.034 | -5.2488
2.808 | -34.0343 | -5.2599
2.8141 | -34.0346 | -5.2713
2.8202 | -34.0349 | -5.2824
2.8263 | -34.0353 | -5.2936
2.8324 | -34.0356 | -5.3052
2.8385 | -34.036 | -5.3163
2.8446 | -34.0364 | -5.3273
2.8507 | -34.0368 | -5.3384
2.8568 | -34.0371 | -5.3491
2.8629 | -34.0374 | -5.36
2.869 | -34.0377 | -5.3713
2.8751 | -34.0381 | -5.3823
2.8812 | -34.0385 | -5.3938
2.8873 | -34.0388 | -5.405
2.8934 | -34.0391 | -5.4162
2.8995 | -34.0395 | -5.4276
2.9056 | -34.0398 | -5.4389
2.9117 | -34.0401 | -5.4502
2.9178 | -34.0405 | -5.4611
2.9239 | -34.0409 | -5.4729
2.93 | -34.0413 | -5.4845
2.9361 | -34.0416 | -5.4956
2.9423 | -34.042 | -5.5065
2.9484 | -34.0424 | -5.5178
2.9545 | -34.0428 | -5.529
2.9606 | -34.0432 | -5.5406
2.9667 | -34.0435 | -5.5519
2.9728 | -34.0439 | -5.5634
2.9789 | -34.0443 | -5.5747
2.985 | -34.0447 | -5.5861
2.9911 | -34.0451 | -5.5976
2.9972 | -34.0456 | -5.6089
3.0033 | -34.046 | -5.6205
3.0094 | -34.0464 | -5.6318
3.0155 | -34.0468 | -5.6429
3.0216 | -34.0472 | -5.6544
3.0277 | -34.0475 | -5.6662
3.0338 | -34.0479 | -5.6778
3.0399 | -34.0483 | -5.6893
3.046 | -34.0486 | -5.701
3.0521 | -34.049 | -5.7123
3.0582 | -34.0493 | -5.7244
3.0643 | -34.0496 | -5.736
3.0704 | -34.05 | -5.7479
3.0765 | -34.0502 | -5.7592
3.0827 | -34.0505 | -5.7706
3.0888 | -34.0509 | -5.782
3.0949 | -34.0513 | -5.7935
3.101 | -34.0516 | -5.8051
3.1071 | -34.052 | -5.8169
3.1132 | -34.0523 | -5.8288
3.1193 | -34.0527 | -5.8404
3.1254 | -34.0531 | -5.8519
3.1315 | -34.0535 | -5.8639
3.1376 | -34.0537 | -5.8754
3.1437 | -34.0541 | -5.8872
3.1498 | -34.0544 | -5.8993
3.1559 | -34.0548 | -5.9109
3.162 | -34.055 | -5.923
3.1681 | -34.0553 | -5.9351
3.1742 | -34.0556 | -5.9467
3.1803 | -34.0559 | -5.9585
3.1864 | -34.0563 | -5.9698
3.1925 | -34.0566 | -5.9815
3.1986 | -34.0569 | -5.9935
3.2047 | -34.0571 | -6.0056
3.2108 | -34.0574 | -6.0169
3.2169 | -34.0577 | -6.0286
3.223 | -34.058 | -6.0404
3.2292 | -34.0582 | -6.0523
3.2353 | -34.0584 | -6.0633
3.2414 | -34.0587 | -6.0748
3.2475 | -34.0589 | -6.0864
3.2536 | -34.0592 | -6.0979
3.2597 | -34.0595 | -6.109
3.2658 | -34.0597 | -6.1201
3.2719 | -34.0599 | -6.1313
3.278 | -34.0602 | -6.143
3.2841 | -34.0604 | -6.1542
3.2902 | -34.0607 | -6.1657
3.2963 | -34.061 | -6.1773
3.3024 | -34.0611 | -6.1885
3.3085 | -34.0614 | -6.2002
3.3146 | -34.0615 | -6.2115
3.3207 | -34.0618 | -6.2231
3.3268 | -34.062 | -6.2347
3.3329 | -34.0622 | -6.2455
3.339 | -34.0624 | -6.2564
3.3451 | -34.0626 | -6.2678
3.3512 | -34.0628 | -6.2786
3.3573 | -34.0631 | -6.2899
3.3634 | -34.0633 | -6.3012
3.3696 | -34.0635 | -6.3127
3.3757 | -34.0636 | -6.3236
3.3818 | -34.0639 | -6.3347
3.3879 | -34.0641 | -6.3459
3.394 | -34.0642 | -6.3574
3.4001 | -34.0644 | -6.3684
3.4062 | -34.0645 | -6.3795
3.4123 | -34.0647 | -6.3902
3.4184 | -34.0649 | -6.4017
3.4245 | -34.0652 | -6.4126
3.4306 | -34.0653 | -6.4235
3.4367 | -34.0655 | -6.4342
3.4428 | -34.0657 | -6.4448
3.4489 | -34.0658 | -6.4559
3.455 | -34.066 | -6.4669
3.4611 | -34.0662 | -6.4783
3.4672 | -34.0664 | -6.4895
3.4733 | -34.0666 | -6.5008
3.4794 | -34.0668 | -6.5118
3.4855 | -34.0669 | -6.5224
3.4916 | -34.0671 | -6.533
3.4977 | -34.0673 | -6.5444
3.5038 | -34.0675 | -6.5555
3.5099 | -34.0676 | -6.5659
3.5161 | -34.0678 | -6.5766
3.5222 | -34.068 | -6.5872
3.5283 | -34.0681 | -6.5982
3.5344 | -34.0682 | -6.6092
3.5405 | -34.0684 | -6.6199
3.5466 | -34.0686 | -6.631
3.5527 | -34.0687 | -6.6421
3.5588 | -34.0689 | -6.6529
3.5649 | -34.069 | -6.6639
3.571 | -34.0691 | -6.6747
3.5771 | -34.0693 | -6.686
3.5832 | -34.0694 | -6.6961
3.5893 | -34.0697 | -6.7069
3.5954 | -34.0699 | -6.718
3.6015 | -34.0701 | -6.7291
3.6076 | -34.0703 | -6.7398
3.6137 | -34.0704 | -6.7502
3.6198 | -34.0707 | -6.7607
3.6259 | -34.0709 | -6.7719
3.632 | -34.0711 | -6.7827
3.6381 | -34.0713 | -6.7937
3.6442 | -34.0714 | -6.8045
3.6503 | -34.0717 | -6.8149
3.6565 | -34.0719 | -6.8259
3.6626 | -34.0721 | -6.8367
3.6687 | -34.0723 | -6.8476
3.6748 | -34.0726 | -6.8582
3.6809 | -34.0729 | -6.8692
3.687 | -34.073 | -6.8804
3.6931 | -34.0733 | -6.8908
3.6992 | -34.0735 | -6.9018
3.7053 | -34.0738 | -6.9125
3.7114 | -34.0739 | -6.9235
3.7175 | -34.0742 | -6.9344
3.7236 | -34.0744 | -6.9447
3.7297 | -34.0747 | -6.9557
3.7358 | -34.0749 | -6.9668
3.7419 | -34.0751 | -6.9779
3.748 | -34.0753 | -6.9886
3.7541 | -34.0756 | -6.9995
3.7602 | -34.0759 | -7.0103
3.7663 | -34.0761 | -7.021
3.7724 | -34.0763 | -7.0319
3.7785 | -34.0766 | -7.0426
3.7846 | -34.0768 | -7.0533
3.7907 | -34.0771 | -7.0645
3.7969 | -34.0774 | -7.0754
3.803 | -34.0777 | -7.0859
3.8091 | -34.0779 | -7.0967
3.8152 | -34.0782 | -7.1075
3.8213 | -34.0784 | -7.1186
3.8274 | -34.0787 | -7.1291
3.8335 | -34.0791 | -7.1401
3.8396 | -34.0794 | -7.151
3.8457 | -34.0797 | -7.162
3.8518 | -34.08 | -7.1728
3.8579 | -34.0804 | -7.1836
3.864 | -34.0806 | -7.1946
3.8701 | -34.0809 | -7.2059
3.8762 | -34.0812 | -7.2165
3.8823 | -34.0814 | -7.2276
3.8884 | -34.0818 | -7.2385
3.8945 | -34.0821 | -7.2496
3.9006 | -34.0824 | -7.2608
3.9067 | -34.0827 | -7.2722
3.9128 | -34.083 | -7.2831
3.9189 | -34.0833 | -7.2939
3.925 | -34.0835 | -7.3049
3.9311 | -34.0838 | -7.3154
3.9372 | -34.0841 | -7.3265
3.9434 | -34.0844 | -7.3375
3.9495 | -34.0848 | -7.3485
3.9556 | -34.0851 | -7.3592
3.9617 | -34.0854 | -7.3697
3.9678 | -34.0858 | -7.3804
3.9739 | -34.086 | -7.3913
3.98 | -34.0863 | -7.4023
3.9861 | -34.0866 | -7.4132
3.9922 | -34.0868 | -7.4241
3.9983 | -34.0871 | -7.4346
4.0044 | -34.0874 | -7.4451
4.0105 | -34.0877 | -7.456
4.0166 | -34.088 | -7.467
4.0227 | -34.0883 | -7.4778
4.0288 | -34.0886 | -7.4891
4.0349 | -34.0888 | -7.4998
4.041 | -34.0891 | -7.5106
4.0471 | -34.0893 | -7.5216
4.0532 | -34.0896 | -7.5321
4.0593 | -34.0899 | -7.5432
4.0654 | -34.0902 | -7.5538
4.0715 | -34.0905 | -7.5651
4.0776 | -34.0908 | -7.5759
4.0838 | -34.0911 | -7.5872
4.0899 | -34.0914 | -7.5979
4.096 | -34.0916 | -7.6087
4.1021 | -34.0919 | -7.6193
4.1082 | -34.0922 | -7.63
4.1143 | -34.0924 | -7.6404
4.1204 | -34.0927 | -7.651
4.1265 | -34.093 | -7.6616
4.1326 | -34.0933 | -7.6722
4.1387 | -34.0936 | -7.6829
4.1448 | -34.0938 | -7.6938
4.1509 | -34.094 | -7.7046
4.157 | -34.0943 | -7.7158
4.1631 | -34.0946 | -7.7268
4.1692 | -34.0949 | -7.7375
4.1753 | -34.0951 | -7.7482
4.1814 | -34.0953 | -7.7587
4.1875 | -34.0956 | -7.7697
4.1936 | -34.0959 | -7.7806
4.1997 | -34.0961 | -7.791
4.2058 | -34.0964 | -7.802
4.2119 | -34.0966 | -7.8126
4.218 | -34.0968 | -7.8231
4.2241 | -34.0971 | -7.8337
4.2303 | -34.0974 | -7.8443
4.2364 | -34.0976 | -7.855
4.2425 | -34.0978 | -7.8657
4.2486 | -34.098 | -7.8766
4.2547 | -34.0982 | -7.8871
4.2608 | -34.0984 | -7.8979
4.2669 | -34.0986 | -7.9083
4.273 | -34.0988 | -7.9191
4.2791 | -34.0991 | -7.9298
4.2852 | -34.0993 | -7.9407
4.2913 | -34.0995 | -7.9511
4.2974 | -34.0997 | -7.9621
4.3035 | -34.1 | -7.973
4.3096 | -34.1001 | -7.9839
4.3157 | -34.1004 | -7.9944
4.3218 | -34.1006 | -8.0049
4.3279 | -34.1008 | -8.0156
4.334 | -34.101 | -8.0262
4.3401 | -34.1012 | -8.0365
4.3462 | -34.1015 | -8.0473
4.3523 | -34.1017 | -8.0575
4.3584 | -34.1019 | -8.0677
4.3645 | -34.1021 | -8.0783
4.3707 | -34.1023 | -8.0887
4.3768 | -34.1025 | -8.0992
4.3829 | -34.1028 | -8.1101
4.389 | -34.103 | -8.1207
4.3951 | -34.1032 | -8.131
4.4012 | -34.1034 | -8.1419
4.4073 | -34.1036 | -8.1523
4.4134 | -34.1038 | -8.1628
4.4195 | -34.1041 | -8.1734
4.4256 | -34.1043 | -8.1837
4.4317 | -34.1045 | -8.1942
4.4378 | -34.1047 | -8.2044
4.4439 | -34.1049 | -8.2147
4.45 | -34.1051 | -8.2252
4.4561 | -34.1052 | -8.2357
4.4622 | -34.1054 | -8.2462
4.4683 | -34.1057 | -8.2567
4.4744 | -34.1059 | -8.2675
4.4805 | -34.1061 | -8.2776
4.4866 | -34.1063 | -8.2881
4.4927 | -34.1065 | -8.2988
4.4988 | -34.1067 | -8.3092
4.5049 | -34.107 | -8.3194
4.511 | -34.1072 | -8.3296
4.5172 | -34.1074 | -8.34
4.5233 | -34.1076 | -8.3509
4.5294 | -34.1079 | -8.3612
4.5355 | -34.1081 | -8.3717
4.5416 | -34.1083 | -8.3821
4.5477 | -34.1085 | -8.3925
4.5538 | -34.1087 | -8.4026
4.5599 | -34.109 | -8.413
4.566 | -34.1092 | -8.4235
4.5721 | -34.1094 | -8.4339
4.5782 | -34.1097 | -8.4446
4.5843 | -34.1099 | -8.4551
4.5904 | -34.1102 | -8.4656
4.5965 | -34.1104 | -8.4759
4.6026 | -34.1107 | -8.4867
4.6087 | -34.1109 | -8.4971
4.6148 | -34.1111 | -8.5078
4.6209 | -34.1113 | -8.5182
4.627 | -34.1116 | -8.5288
4.6331 | -34.1118 | -8.5392
4.6392 | -34.112 | -8.5497
4.6453 | -34.1122 | -8.5601
4.6514 | -34.1125 | -8.5706
4.6576 | -34.1127 | -8.5814
4.6637 | -34.1129 | -8.5915
4.6698 | -34.1131 | -8.6018
4.6759 | -34.1133 | -8.6127
4.682 | -34.1135 | -8.623
4.6881 | -34.1137 | -8.6335
4.6942 | -34.114 | -8.6439
4.7003 | -34.1142 | -8.6544
4.7064 | -34.1144 | -8.6647
4.7125 | -34.1147 | -8.6755
4.7186 | -34.1149 | -8.6857
4.7247 | -34.1152 | -8.6963
4.7308 | -34.1154 | -8.7067
4.7369 | -34.1156 | -8.7173
4.743 | -34.1159 | -8.7281
4.7491 | -34.1161 | -8.7389
4.7552 | -34.1163 | -8.7495
4.7613 | -34.1166 | -8.7601
4.7674 | -34.1168 | -8.7705
4.7735 | -34.117 | -8.7814
4.7796 | -34.1173 | -8.7921
4.7857 | -34.1175 | -8.8027
4.7918 | -34.1178 | -8.813
4.7979 | -34.1181 | -8.8235
4.8041 | -34.1183 | -8.834
4.8102 | -34.1186 | -8.8442
4.8163 | -34.1188 | -8.8547
4.8224 | -34.119 | -8.8652
4.8285 | -34.1193 | -8.8754
4.8346 | -34.1196 | -8.8858
4.8407 | -34.1198 | -8.8964
4.8468 | -34.1201 | -8.9072
4.8529 | -34.1203 | -8.918
4.859 | -34.1206 | -8.9288
4.8651 | -34.1208 | -8.9392
4.8712 | -34.1211 | -8.9495
4.8773 | -34.1213 | -8.9601
4.8834 | -34.1216 | -8.9707
4.8895 | -34.1218 | -8.981
4.8956 | -34.122 | -8.9915
4.9017 | -34.1222 | -9.002
4.9078 | -34.1225 | -9.0128
4.9139 | -34.1228 | -9.023
4.92 | -34.1231 | -9.0332
4.9261 | -34.1234 | -9.0434
4.9322 | -34.1236 | -9.054
4.9383 | -34.1239 | -9.0645
4.9445 | -34.1242 | -9.075
4.9506 | -34.1244 | -9.0854
4.9567 | -34.1247 | -9.0957
4.9628 | -34.1249 | -9.1062
4.9689 | -34.1252 | -9.1171
4.975 | -34.1254 | -9.1279
4.9811 | -34.1257 | -9.1382
4.9872 | -34.126 | -9.1485
4.9933 | -34.1264 | -9.1593
4.9994 | -34.1266 | -9.1703
5.0055 | -34.1269 | -9.1808
5.0116 | -34.1272 | -9.1917
5.0177 | -34.1275 | -9.2025
5.0238 | -34.1278 | -9.2134
5.0299 | -34.1282 | -9.2243
5.036 | -34.1285 | -9.2348
5.0421 | -34.1287 | -9.2453
5.0482 | -34.129 | -9.2559
5.0543 | -34.1293 | -9.2662
5.0604 | -34.1296 | -9.2768
5.0665 | -34.1299 | -9.2871
5.0726 | -34.1302 | -9.298
5.0787 | -34.1305 | -9.3088
5.0848 | -34.1308 | -9.3195
5.091 | -34.1311 | -9.3299
5.0971 | -34.1314 | -9.3402
5.1032 | -34.1317 | -9.3511
5.1093 | -34.132 | -9.3617
5.1154 | -34.1323 | -9.3721
5.1215 | -34.1325 | -9.3825
5.1276 | -34.1328 | -9.3929
5.1337 | -34.1332 | -9.4033
5.1398 | -34.1335 | -9.4136
5.1459 | -34.1338 | -9.4245
5.152 | -34.1341 | -9.4352
5.1581 | -34.1344 | -9.4459
5.1642 | -34.1347 | -9.4566
5.1703 | -34.135 | -9.4671
5.1764 | -34.1353 | -9.4775
5.1825 | -34.1356 | -9.4879
5.1886 | -34.1359 | -9.4981
5.1947 | -34.1362 | -9.5087
5.2008 | -34.1365 | -9.5194
5.2069 | -34.1368 | -9.5301
5.213 | -34.1371 | -9.5405
5.2191 | -34.1375 | -9.5512
5.2252 | -34.1377 | -9.5618
5.2314 | -34.138 | -9.5726
5.2375 | -34.1383 | -9.5829
5.2436 | -34.1386 | -9.5935
5.2497 | -34.1389 | -9.6038
5.2558 | -34.1393 | -9.6147
5.2619 | -34.1396 | -9.6256
5.268 | -34.1399 | -9.6363
5.2741 | -34.1403 | -9.647
5.2802 | -34.1405 | -9.6579
5.2863 | -34.1409 | -9.6684
5.2924 | -34.1412 | -9.6792
5.2985 | -34.1415 | -9.6899
5.3046 | -34.1419 | -9.7006
5.3107 | -34.1422 | -9.7114
5.3168 | -34.1425 | -9.722
5.3229 | -34.1428 | -9.7327
5.329 | -34.1431 | -9.7433
5.3351 | -34.1434 | -9.7536
5.3412 | -34.1437 | -9.7642
5.3473 | -34.144 | -9.7749
5.3534 | -34.1444 | -9.7856
5.3595 | -34.1447 | -9.7963
5.3656 | -34.145 | -9.8068
5.3717 | -34.1453 | -9.8177
5.3779 | -34.1456 | -9.8287
5.384 | -34.1459 | -9.8393
5.3901 | -34.1463 | -9.8502
5.3962 | -34.1466 | -9.8607
5.4023 | -34.1469 | -9.8715
5.4084 | -34.1473 | -9.882
5.4145 | -34.1476 | -9.8924
5.4206 | -34.1479 | -9.9031
5.4267 | -34.1483 | -9.9141
5.4328 | -34.1486 | -9.9247
5.4389 | -34.1489 | -9.9351
5.445 | -34.1492 | -9.946
5.4511 | -34.1495 | -9.9565
5.4572 | -34.1499 | -9.9672
5.4633 | -34.1502 | -9.9779
5.4694 | -34.1505 | -9.9885
5.4755 | -34.1508 | -9.9989
5.4816 | -34.1511 | -10.0096
5.4877 | -34.1514 | -10.0201
5.4938 | -34.1517 | -10.0305
5.4999 | -34.1521 | -10.041
5.506 | -34.1524 | -10.0516
5.5121 | -34.1527 | -10.0621
5.5183 | -34.1531 | -10.0724
5.5244 | -34.1533 | -10.0833
5.5305 | -34.1537 | -10.0938
5.5366 | -34.154 | -10.1042
5.5427 | -34.1543 | -10.1149
5.5488 | -34.1546 | -10.1256
5.5549 | -34.1549 | -10.1362
5.561 | -34.1552 | -10.147
5.5671 | -34.1556 | -10.1574
5.5732 | -34.1559 | -10.1683
5.5793 | -34.1562 | -10.1791
5.5854 | -34.1566 | -10.1897
5.5915 | -34.1569 | -10.2006
5.5976 | -34.1572 | -10.2112
5.6037 | -34.1575 | -10.2219
5.6098 | -34.1578 | -10.2325
5.6159 | -34.1581 | -10.2429
5.622 | -34.1584 | -10.2536
5.6281 | -34.1588 | -10.2642
5.6342 | -34.1591 | -10.2751
5.6403 | -34.1594 | -10.2858
5.6464 | -34.1597 | -10.2966
5.6525 | -34.16 | -10.3073
5.6586 | -34.1603 | -10.3176
5.6648 | -34.1607 | -10.3283
5.6709 | -34.161 | -10.3391
5.677 | -34.1614 | -10.3497
5.6831 | -34.1617 | -10.3604
5.6892 | -34.162 | -10.3712
5.6953 | -34.1623 | -10.3817
5.7014 | -34.1626 | -10.3924
5.7075 | -34.1629 | -10.403
5.7136 | -34.1632 | -10.4136
5.7197 | -34.1635 | -10.4242
5.7258 | -34.1638 | -10.4346
5.7319 | -34.1641 | -10.4452
5.738 | -34.1644 | -10.4558
5.7441 | -34.1648 | -10.4664
5.7502 | -34.1651 | -10.4768
5.7563 | -34.1654 | -10.4874
5.7624 | -34.1658 | -10.4982
5.7685 | -34.1661 | -10.5088
5.7746 | -34.1664 | -10.5196
5.7807 | -34.1667 | -10.5303
5.7868 | -34.1669 | -10.5411
5.7929 | -34.1672 | -10.5516
5.799 | -34.1676 | -10.562
5.8052 | -34.1679 | -10.5725
5.8113 | -34.1682 | -10.583
5.8174 | -34.1685 | -10.5934
5.8235 | -34.1688 | -10.604
5.8296 | -34.1691 | -10.6147
5.8357 | -34.1694 | -10.6252
5.8418 | -34.1697 | -10.6359
5.8479 | -34.17 | -10.6466
5.854 | -34.1703 | -10.6573
5.8601 | -34.1706 | -10.6678
5.8662 | -34.1709 | -10.6786
5.8723 | -34.1712 | -10.689
5.8784 | -34.1715 | -10.6994
5.8845 | -34.1718 | -10.7099
5.8906 | -34.1721 | -10.7206
5.8967 | -34.1724 | -10.731
5.9028 | -34.1728 | -10.7414
5.9089 | -34.1732 | -10.7518
5.915 | -34.1734 | -10.7621
5.9211 | -34.1737 | -10.7726
5.9272 | -34.1741 | -10.7829
5.9333 | -34.1744 | -10.7933
5.9394 | -34.1747 | -10.8038
5.9455 | -34.175 | -10.8141
5.9517 | -34.1754 | -10.8245
5.9578 | -34.1757 | -10.8352
5.9639 | -34.1761 | -10.8454
5.97 | -34.1764 | -10.856
5.9761 | -34.1768 | -10.8664
5.9822 | -34.177 | -10.8771
5.9883 | -34.1774 | -10.8876
5.9944 | -34.1777 | -10.8982
6.0005 | -34.178 | -10.9089
6.0066 | -34.1783 | -10.9192
6.0127 | -34.1787 | -10.9294
6.0188 | -34.179 | -10.9401
6.0249 | -34.1793 | -10.9507
6.031 | -34.1796 | -10.961
6.0371 | -34.1799 | -10.9715
6.0432 | -34.1803 | -10.9819
6.0493 | -34.1806 | -10.9923
6.0554 | -34.1809 | -11.0026
6.0615 | -34.1813 | -11.0132
6.0676 | -34.1816 | -11.0238
6.0737 | -34.1819 | -11.0344
6.0798 | -34.1822 | -11.0449
6.0859 | -34.1826 | -11.0553
6.0921 | -34.1829 | -11.0656
6.0982 | -34.1832 | -11.076
6.1043 | -34.1836 | -11.0864
6.1104 | -34.1839 | -11.0971
6.1165 | -34.1842 | -11.1071
6.1226 | -34.1846 | -11.1175
6.1287 | -34.1849 | -11.1279
6.1348 | -34.1852 | -11.1384
6.1409 | -34.1855 | -11.1488
6.147 | -34.1859 | -11.1592
6.1531 | -34.1862 | -11.1696
6.1592 | -34.1865 | -11.1799
6.1653 | -34.1868 | -11.1902
6.1714 | -34.1871 | -11.2007
6.1775 | -34.1875 | -11.2112
6.1836 | -34.1878 | -11.2217
6.1897 | -34.1881 | -11.2321
6.1958 | -34.1884 | -11.2428
6.2019 | -34.1888 | -11.2533
6.208 | -34.1891 | -11.2637
6.2141 | -34.1895 | -11.2741
6.2202 | -34.1898 | -11.2846
6.2263 | -34.1902 | -11.295
6.2325 | -34.1905 | -11.3052
6.2386 | -34.1908 | -11.3153
6.2447 | -34.1911 | -11.326
6.2508 | -34.1915 | -11.3364
6.2569 | -34.1918 | -11.3468
6.263 | -34.1921 | -11.3572
6.2691 | -34.1924 | -11.3675
6.2752 | -34.1927 | -11.378
6.2813 | -34.1931 | -11.3884
6.2874 | -34.1934 | -11.3987
6.2935 | -34.1937 | -11.409
6.2996 | -34.194 | -11.4194
6.3057 | -34.1944 | -11.4298
6.3118 | -34.1946 | -11.4403
6.3179 | -34.195 | -11.4507
6.324 | -34.1953 | -11.4612
6.3301 | -34.1956 | -11.4716
6.3362 | -34.196 | -11.4819
6.3423 | -34.1963 | -11.4925
6.3484 | -34.1967 | -11.503
6.3545 | -34.197 | -11.5131
6.3606 | -34.1973 | -11.5232
6.3667 | -34.1976 | -11.5335
6.3728 | -34.1979 | -11.5439
6.379 | -34.1982 | -11.5543
6.3851 | -34.1986 | -11.5645
6.3912 | -34.1989 | -11.5749
6.3973 | -34.1992 | -11.5853
6.4034 | -34.1995 | -11.5955
6.4095 | -34.1998 | -11.6059
6.4156 | -34.2001 | -11.6163
6.4217 | -34.2004 | -11.6266
6.4278 | -34.2008 | -11.637
6.4339 | -34.2011 | -11.6473
6.44 | -34.2015 | -11.6576
6.4461 | -34.2018 | -11.6681
6.4522 | -34.2021 | -11.6783
6.4583 | -34.2024 | -11.6888
6.4644 | -34.2027 | -11.6992
6.4705 | -34.203 | -11.7093
6.4766 | -34.2033 | -11.7197
6.4827 | -34.2037 | -11.7301
6.4888 | -34.2039 | -11.7403
6.4949 | -34.2043 | -11.7504
6.501 | -34.2046 | -11.7609
6.5071 | -34.2049 | -11.7711
6.5132 | -34.2053 | -11.7814
6.5194 | -34.2056 | -11.7918
6.5255 | -34.2059 | -11.802
6.5316 | -34.2062 | -11.8124
6.5377 | -34.2065 | -11.8228
6.5438 | -34.2068 | -11.833
6.5499 | -34.2072 | -11.8432
6.556 | -34.2075 | -11.8535
6.5621 | -34.2078 | -11.864
6.5682 | -34.2081 | -11.8744
6.5743 | -34.2085 | -11.8847
6.5804 | -34.2088 | -11.8952
6.5865 | -34.2091 | -11.9055
6.5926 | -34.2095 | -11.9158
6.5987 | -34.2098 | -11.9259
6.6048 | -34.2101 | -11.9363
6.6109 | -34.2104 | -11.9467
6.617 | -34.2107 | -11.9571
6.6231 | -34.211 | -11.9674
6.6292 | -34.2114 | -11.9778
6.6353 | -34.2117 | -11.9884
6.6414 | -34.2121 | -11.9986
6.6475 | -34.2124 | -12.009
6.6536 | -34.2127 | -12.0193
6.6597 | -34.213 | -12.0294
6.6659 | -34.2133 | -12.0397
6.672 | -34.2137 | -12.05
6.6781 | -34.214 | -12.0603
6.6842 | -34.2143 | -12.0705
6.6903 | -34.2147 | -12.0808
6.6964 | -34.215 | -12.0911
6.7025 | -34.2153 | -12.1012
6.7086 | -34.2156 | -12.1115
6.7147 | -34.216 | -12.1219
6.7208 | -34.2163 | -12.1323
6.7269 | -34.2166 | -12.1426
6.733 | -34.2169 | -12.1527
6.7391 | -34.2172 | -12.163
6.7452 | -34.2176 | -12.1731
6.7513 | -34.2179 | -12.1832
6.7574 | -34.2182 | -12.1936
6.7635 | -34.2186 | -12.2037
6.7696 | -34.2189 | -12.2139
6.7757 | -34.2193 | -12.2242
6.7818 | -34.2196 | -12.2345
6.7879 | -34.2199 | -12.2446
6.794 | -34.2203 | -12.2547
6.8001 | -34.2205 | -12.265
6.8063 | -34.2209 | -12.2753
6.8124 | -34.2212 | -12.2857
6.8185 | -34.2215 | -12.2962
6.8246 | -34.2218 | -12.3064
6.8307 | -34.2221 | -12.3166
6.8368 | -34.2225 | -12.327
6.8429 | -34.2228 | -12.3374
6.849 | -34.2231 | -12.3477
6.8551 | -34.2235 | -12.3578
6.8612 | -34.2238 | -12.3681
6.8673 | -34.2241 | -12.3783
6.8734 | -34.2245 | -12.3888
6.8795 | -34.2248 | -12.399
6.8856 | -34.2251 | -12.4093
6.8917 | -34.2254 | -12.4196
6.8978 | -34.2258 | -12.4296
6.9039 | -34.2261 | -12.4396
6.91 | -34.2265 | -12.4499
6.9161 | -34.2268 | -12.46
6.9222 | -34.2271 | -12.4702
6.9283 | -34.2275 | -12.4802
6.9344 | -34.2278 | -12.4905
6.9405 | -34.2282 | -12.5006
6.9466 | -34.2285 | -12.5111
6.9528 | -34.2289 | -12.5215
6.9589 | -34.2292 | -12.5319
6.965 | -34.2296 | -12.5422
6.9711 | -34.2299 | -12.5526
6.9772 | -34.2303 | -12.5628
6.9833 | -34.2306 | -12.573
6.9894 | -34.2309 | -12.5834
6.9955 | -34.2313 | -12.5935
7.0016 | -34.2316 | -12.6037
7.0077 | -34.232 | -12.6139
7.0138 | -34.2323 | -12.6241
7.0199 | -34.2326 | -12.6343
7.026 | -34.233 | -12.6445
7.0321 | -34.2333 | -12.6545
7.0382 | -34.2337 | -12.6645
7.0443 | -34.234 | -12.6748
7.0504 | -34.2344 | -12.6851
7.0565 | -34.2347 | -12.6953
7.0626 | -34.235 | -12.7053
7.0687 | -34.2354 | -12.7151
7.0748 | -34.2357 | -12.7251
7.0809 | -34.2361 | -12.7353
7.087 | -34.2364 | -12.7456
7.0932 | -34.2367 | -12.7558
7.0993 | -34.2371 | -12.766
7.1054 | -34.2374 | -12.776
7.1115 | -34.2378 | -12.786
7.1176 | -34.2381 | -12.7961
7.1237 | -34.2384 | -12.8062
7.1298 | -34.2388 | -12.8163
7.1359 | -34.2391 | -12.8265
7.142 | -34.2395 | -12.8364
7.1481 | -34.2398 | -12.8464
7.1542 | -34.2401 | -12.8566
7.1603 | -34.2405 | -12.8668
7.1664 | -34.2409 | -12.8769
7.1725 | -34.2412 | -12.887
7.1786 | -34.2416 | -12.8971
7.1847 | -34.2419 | -12.9073
7.1908 | -34.2422 | -12.9176
7.1969 | -34.2426 | -12.9279
7.203 | -34.2429 | -12.9378
7.2091 | -34.2433 | -12.9478
7.2152 | -34.2436 | -12.958
7.2213 | -34.244 | -12.9679
7.2274 | -34.2443 | -12.978
7.2335 | -34.2447 | -12.9881
7.2397 | -34.245 | -12.9981
7.2458 | -34.2454 | -13.008
7.2519 | -34.2458 | -13.0181
7.258 | -34.2461 | -13.0284
7.2641 | -34.2464 | -13.0385
7.2702 | -34.2468 | -13.0488
7.2763 | -34.2471 | -13.0589
7.2824 | -34.2475 | -13.0691
7.2885 | -34.2479 | -13.0792
7.2946 | -34.2482 | -13.0893
7.3007 | -34.2485 | -13.0995
7.3068 | -34.2488 | -13.1097
7.3129 | -34.2492 | -13.1199
7.319 | -34.2496 | -13.13
7.3251 | -34.2499 | -13.1399
7.3312 | -34.2503 | -13.15
7.3373 | -34.2506 | -13.16
7.3434 | -34.251 | -13.17
7.3495 | -34.2514 | -13.1799
7.3556 | -34.2517 | -13.1901
7.3617 | -34.2521 | -13.2003
7.3678 | -34.2525 | -13.2103
7.3739 | -34.2528 | -13.2205
7.3801 | -34.2531 | -13.2309
7.3862 | -34.2535 | -13.2408
7.3923 | -34.2539 | -13.2511
7.3984 | -34.2542 | -13.261
7.4045 | -34.2546 | -13.271
7.4106 | -34.255 | -13.281
7.4167 | -34.2553 | -13.2912
7.4228 | -34.2557 | -13.3014
7.4289 | -34.256 | -13.3114
7.435 | -34.2564 | -13.3213
7.4411 | -34.2567 | -13.3314
7.4472 | -34.2571 | -13.3415
7.4533 | -34.2575 | -13.3513
7.4594 | -34.2578 | -13.3616
7.4655 | -34.2582 | -13.3715
7.4716 | -34.2585 | -13.3818
7.4777 | -34.2589 | -13.3918
7.4838 | -34.2592 | -13.4017
7.4899 | -34.2596 | -13.4116
7.496 | -34.2599 | -13.4215
7.5021 | -34.2603 | -13.4315
7.5082 | -34.2607 | -13.4416
7.5143 | -34.261 | -13.4517
7.5204 | -34.2614 | -13.4619
7.5266 | -34.2617 | -13.4717
7.5327 | -34.2621 | -13.4817
7.5388 | -34.2624 | -13.492
7.5449 | -34.2628 | -13.5022
7.551 | -34.2632 | -13.5121
7.5571 | -34.2635 | -13.5223
7.5632 | -34.2639 | -13.5324
7.5693 | -34.2643 | -13.5423
7.5754 | -34.2646 | -13.5522
7.5815 | -34.265 | -13.5622
7.5876 | -34.2653 | -13.5725
7.5937 | -34.2657 | -13.5826
7.5998 | -34.2661 | -13.5926
7.6059 | -34.2665 | -13.6029
7.612 | -34.2668 | -13.6129
7.6181 | -34.2672 | -13.6228
7.6242 | -34.2676 | -13.6332
7.6303 | -34.2679 | -13.6433
7.6364 | -34.2683 | -13.6535
7.6425 | -34.2686 | -13.6636
7.6486 | -34.269 | -13.6737
7.6547 | -34.2694 | -13.6836
7.6608 | -34.2697 | -13.6936
7.667 | -34.2701 | -13.7037
7.6731 | -34.2704 | -13.7138
7.6792 | -34.2708 | -13.7241
7.6853 | -34.2711 | -13.7343
7.6914 | -34.2715 | -13.7444
7.6975 | -34.2719 | -13.7545
7.7036 | -34.2723 | -13.7646
7.7097 | -34.2726 | -13.7748
7.7158 | -34.273 | -13.7849
7.7219 | -34.2734 | -13.7948
7.728 | -34.2737 | -13.805
7.7341 | -34.2741 | -13.8151
7.7402 | -34.2745 | -13.8252
7.7463 | -34.2749 | -13.8353
7.7524 | -34.2752 | -13.8457
7.7585 | -34.2756 | -13.8557
7.7646 | -34.2759 | -13.8658
7.7707 | -34.2763 | -13.876
7.7768 | -34.2766 | -13.8862
7.7829 | -34.277 | -13.8965
7.789 | -34.2773 | -13.9068
7.7951 | -34.2777 | -13.9169
7.8012 | -34.2781 | -13.927
7.8073 | -34.2785 | -13.9372
7.8135 | -34.2788 | -13.9472
7.8196 | -34.2792 | -13.9575
7.8257 | -34.2796 | -13.9678
7.8318 | -34.2799 | -13.978
7.8379 | -34.2803 | -13.988
7.844 | -34.2807 | -13.998
7.8501 | -34.281 | -14.0083
7.8562 | -34.2814 | -14.0186
7.8623 | -34.2818 | -14.0286
7.8684 | -34.2821 | -14.0388
7.8745 | -34.2825 | -14.049
7.8806 | -34.2828 | -14.0592
7.8867 | -34.2832 | -14.0691
7.8928 | -34.2836 | -14.079
7.8989 | -34.284 | -14.0892
7.905 | -34.2843 | -14.0994
7.9111 | -34.2847 | -14.1095
7.9172 | -34.2851 | -14.1197
7.9233 | -34.2854 | -14.1298
7.9294 | -34.2858 | -14.14
7.9355 | -34.2861 | -14.15
7.9416 | -34.2865 | -14.1602
7.9477 | -34.2868 | -14.1701
7.9539 | -34.2872 | -14.1801
7.96 | -34.2876 | -14.1903
7.9661 | -34.2879 | -14.2004
7.9722 | -34.2883 | -14.2104
7.9783 | -34.2887 | -14.2207
7.9844 | -34.2891 | -14.2308
7.9905 | -34.2894 | -14.2411
7.9966 | -34.2898 | -14.2513
8.0027 | -34.2901 | -14.2614
8.0088 | -34.2905 | -14.2715
8.0149 | -34.2908 | -14.2817
8.021 | -34.2912 | -14.2918
8.0271 | -34.2916 | -14.3021
8.0332 | -34.2919 | -14.3122
8.0393 | -34.2923 | -14.3224
8.0454 | -34.2927 | -14.3327
8.0515 | -34.2931 | -14.3428
8.0576 | -34.2934 | -14.3532
8.0637 | -34.2938 | -14.3632
8.0698 | -34.2941 | -14.3732
8.0759 | -34.2945 | -14.3832
8.082 | -34.2949 | -14.3933
8.0881 | -34.2952 | -14.4035
8.0942 | -34.2956 | -14.4136
8.1004 | -34.296 | -14.4239
8.1065 | -34.2964 | -14.4339
8.1126 | -34.2967 | -14.444
8.1187 | -34.2971 | -14.4543
8.1248 | -34.2974 | -14.4642
8.1309 | -34.2978 | -14.4742
8.137 | -34.2982 | -14.4841
8.1431 | -34.2985 | -14.4945
8.1492 | -34.2989 | -14.5046
8.1553 | -34.2992 | -14.5148
8.1614 | -34.2996 | -14.5251
8.1675 | -34.3 | -14.5352
8.1736 | -34.3003 | -14.5455
8.1797 | -34.3007 | -14.5557
8.1858 | -34.3011 | -14.5659
8.1919 | -34.3014 | -14.576
8.198 | -34.3018 | -14.5859
8.2041 | -34.3022 | -14.5961
8.2102 | -34.3025 | -14.6063
8.2163 | -34.3029 | -14.6162
8.2224 | -34.3033 | -14.6264
8.2285 | -34.3036 | -14.6362
8.2346 | -34.304 | -14.6463
8.2408 | -34.3044 | -14.6563
8.2469 | -34.3047 | -14.6665
8.253 | -34.3051 | -14.6765
8.2591 | -34.3055 | -14.6867
8.2652 | -34.3059 | -14.6968
8.2713 | -34.3062 | -14.707
8.2774 | -34.3066 | -14.7172
8.2835 | -34.307 | -14.7276
8.2896 | -34.3073 | -14.7375
8.2957 | -34.3077 | -14.7477
8.3018 | -34.3081 | -14.7579
8.3079 | -34.3084 | -14.7679
8.314 | -34.3088 | -14.778
8.3201 | -34.3092 | -14.788
8.3262 | -34.3096 | -14.7982
8.3323 | -34.3099 | -14.8083
8.3384 | -34.3103 | -14.8185
8.3445 | -34.3106 | -14.8285
8.3506 | -34.311 | -14.8387
8.3567 | -34.3114 | -14.8488
8.3628 | -34.3118 | -14.859
8.3689 | -34.3121 | -14.869
8.375 | -34.3125 | -14.8791
8.3812 | -34.3128 | -14.889
8.3873 | -34.3132 | -14.8992
8.3934 | -34.3136 | -14.9093
8.3995 | -34.3139 | -14.9193
8.4056 | -34.3143 | -14.9294
8.4117 | -34.3147 | -14.9396
8.4178 | -34.315 | -14.9497
8.4239 | -34.3154 | -14.9597
8.43 | -34.3158 | -14.9698
8.4361 | -34.3161 | -14.98
8.4422 | -34.3165 | -14.99
8.4483 | -34.3169 | -15
8.4544 | -34.3172 | -15.0101
8.4605 | -34.3176 | -15.0201
8.4666 | -34.318 | -15.0303
8.4727 | -34.3183 | -15.0404
8.4788 | -34.3187 | -15.0505
8.4849 | -34.319 | -15.0606
8.491 | -34.3194 | -15.0707
8.4971 | -34.3198 | -15.0806
8.5032 | -34.3201 | -15.0906
8.5093 | -34.3205 | -15.1005
8.5154 | -34.3209 | -15.1106
8.5215 | -34.3213 | -15.1206
8.5277 | -34.3216 | -15.1307
8.5338 | -34.322 | -15.1407
8.5399 | -34.3224 | -15.1508
8.546 | -34.3228 | -15.1607
8.5521 | -34.3231 | -15.1707
8.5582 | -34.3235 | -15.1808
8.5643 | -34.3238 | -15.1906
8.5704 | -34.3242 | -15.2005
8.5765 | -34.3246 | -15.2107
8.5826 | -34.325 | -15.2204
8.5887 | -34.3254 | -15.2304
8.5948 | -34.3258 | -15.2404
8.6009 | -34.3261 | -15.2503
8.607 | -34.3265 | -15.2604
8.6131 | -34.3269 | -15.2705
8.6192 | -34.3273 | -15.2804
8.6253 | -34.3276 | -15.2903
8.6314 | -34.328 | -15.3003
8.6375 | -34.3283 | -15.3104
8.6436 | -34.3287 | -15.3204
8.6497 | -34.3291 | -15.3304
8.6558 | -34.3295 | -15.3404
8.6619 | -34.3298 | -15.3504
8.6681 | -34.3302 | -15.3606
8.6742 | -34.3306 | -15.3706
8.6803 | -34.331 | -15.3805
8.6864 | -34.3314 | -15.3902
8.6925 | -34.3317 | -15.4
8.6986 | -34.3321 | -15.4101
8.7047 | -34.3325 | -15.4203
8.7108 | -34.3328 | -15.4303
8.7169 | -34.3332 | -15.4403
8.723 | -34.3336 | -15.4502
8.7291 | -34.3339 | -15.46
8.7352 | -34.3343 | -15.4702
8.7413 | -34.3347 | -15.4802
8.7474 | -34.335 | -15.4901
8.7535 | -34.3354 | -15.5001
8.7596 | -34.3358 | -15.5101
8.7657 | -34.3362 | -15.5202
8.7718 | -34.3365 | -15.5303
8.7779 | -34.3369 | -15.5404
8.784 | -34.3373 | -15.5504
8.7901 | -34.3377 | -15.5603
8.7962 | -34.3381 | -15.5701
8.8023 | -34.3384 | -15.58
8.8084 | -34.3388 | -15.59
8.8146 | -34.3392 | -15.5999
8.8207 | -34.3396 | -15.6099
8.8268 | -34.34 | -15.6196
8.8329 | -34.3403 | -15.6296
8.839 | -34.3407 | -15.6396
8.8451 | -34.3411 | -15.6494
8.8512 | -34.3415 | -15.6593
8.8573 | -34.3418 | -15.6692
8.8634 | -34.3422 | -15.679
8.8695 | -34.3426 | -15.6891
8.8756 | -34.343 | -15.6991
8.8817 | -34.3434 | -15.709
8.8878 | -34.3437 | -15.7191
8.8939 | -34.3441 | -15.729
8.9 | -34.3445 | -15.7389
8.9061 | -34.3449 | -15.7488
8.9122 | -34.3453 | -15.7588
8.9183 | -34.3457 | -15.7688
8.9244 | -34.3461 | -15.7789
8.9305 | -34.3465 | -15.7886
8.9366 | -34.3469 | -15.7985
8.9427 | -34.3472 | -15.8085
8.9488 | -34.3476 | -15.8184
8.955 | -34.348 | -15.8283
8.9611 | -34.3484 | -15.8384
8.9672 | -34.3488 | -15.8483
8.9733 | -34.3492 | -15.8581
8.9794 | -34.3496 | -15.8679
8.9855 | -34.35 | -15.8779
8.9916 | -34.3504 | -15.888
8.9977 | -34.3508 | -15.8979
9.0038 | -34.3512 | -15.9077
9.0099 | -34.3516 | -15.9176
9.016 | -34.352 | -15.9276
9.0221 | -34.3524 | -15.9373
9.0282 | -34.3528 | -15.947
9.0343 | -34.3531 | -15.9569
9.0404 | -34.3536 | -15.9668
9.0465 | -34.354 | -15.9768
9.0526 | -34.3544 | -15.9867
9.0587 | -34.3548 | -15.9966
9.0648 | -34.3552 | -16.0063
9.0709 | -34.3556 | -16.0162
9.077 | -34.356 | -16.0262
9.0831 | -34.3564 | -16.0362
9.0892 | -34.3568 | -16.0461
9.0953 | -34.3572 | -16.0561
9.1015 | -34.3576 | -16.066
9.1076 | -34.358 | -16.0757
9.1137 | -34.3584 | -16.0856
9.1198 | -34.3588 | -16.0954
9.1259 | -34.3592 | -16.1053
9.132 | -34.3596 | -16.1154
9.1381 | -34.36 | -16.1252
9.1442 | -34.3605 | -16.1348
9.1503 | -34.3608 | -16.1447
9.1564 | -34.3612 | -16.1546
9.1625 | -34.3616 | -16.1645
9.1686 | -34.362 | -16.1743
9.1747 | -34.3625 | -16.1842
9.1808 | -34.3629 | -16.1941
9.1869 | -34.3633 | -16.2039
9.193 | -34.3637 | -16.2138
9.1991 | -34.3641 | -16.2238
9.2052 | -34.3645 | -16.2336
9.2113 | -34.3649 | -16.2435
9.2174 | -34.3654 | -16.2534
9.2235 | -34.3658 | -16.2632
9.2296 | -34.3662 | -16.2731
9.2357 | -34.3666 | -16.2831
9.2419 | -34.367 | -16.2929
9.248 | -34.3675 | -16.3027
9.2541 | -34.3679 | -16.3127
9.2602 | -34.3683 | -16.3227
9.2663 | -34.3687 | -16.3323
9.2724 | -34.3691 | -16.3421
9.2785 | -34.3695 | -16.3519
9.2846 | -34.37 | -16.3617
9.2907 | -34.3704 | -16.3715
9.2968 | -34.3708 | -16.3815
9.3029 | -34.3712 | -16.3914
9.309 | -34.3716 | -16.4013
9.3151 | -34.372 | -16.411
9.3212 | -34.3725 | -16.4207
9.3273 | -34.3729 | -16.4305
9.3334 | -34.3733 | -16.4403
9.3395 | -34.3737 | -16.45
9.3456 | -34.3741 | -16.4597
9.3517 | -34.3745 | -16.4696
9.3578 | -34.3749 | -16.4794
9.3639 | -34.3754 | -16.4891
9.37 | -34.3758 | -16.4989
9.3761 | -34.3762 | -16.5089
9.3822 | -34.3766 | -16.5188
9.3884 | -34.3771 | -16.5284
9.3945 | -34.3775 | -16.5383
9.4006 | -34.3779 | -16.5481
9.4067 | -34.3783 | -16.558
9.4128 | -34.3788 | -16.5678
9.4189 | -34.3792 | -16.5778
9.425 | -34.3796 | -16.5875
9.4311 | -34.38 | -16.5973
9.4372 | -34.3805 | -16.607
9.4433 | -34.3809 | -16.6169
9.4494 | -34.3813 | -16.6266
9.4555 | -34.3817 | -16.6366
9.4616 | -34.3822 | -16.6464
9.4677 | -34.3826 | -16.6563
9.4738 | -34.383 | -16.6661
9.4799 | -34.3835 | -16.6759
9.486 | -34.3839 | -16.6857
9.4921 | -34.3843 | -16.6955
9.4982 | -34.3848 | -16.7051
9.5043 | -34.3852 | -16.7148
9.5104 | -34.3856 | -16.7245
9.5165 | -34.386 | -16.7343
9.5226 | -34.3864 | -16.7443
9.5288 | -34.3869 | -16.754
9.5349 | -34.3873 | -16.7638
9.541 | -34.3877 | -16.7734
9.5471 | -34.3881 | -16.7832
9.5532 | -34.3886 | -16.793
9.5593 | -34.389 | -16.8029
9.5654 | -34.3894 | -16.8127
9.5715 | -34.3898 | -16.8224
9.5776 | -34.3903 | -16.8323
9.5837 | -34.3907 | -16.8421
9.5898 | -34.3911 | -16.8519
9.5959 | -34.3915 | -16.8618
9.602 | -34.3919 | -16.8714
9.6081 | -34.3924 | -16.8811
9.6142 | -34.3928 | -16.8908
9.6203 | -34.3932 | -16.9006
9.6264 | -34.3936 | -16.9103
9.6325 | -34.394 | -16.9202
9.6386 | -34.3945 | -16.9299
9.6447 | -34.3949 | -16.9397
9.6508 | -34.3953 | -16.9494
9.6569 | -34.3957 | -16.9592
9.663 | -34.3962 | -16.969
9.6691 | -34.3966 | -16.9788
9.6753 | -34.3971 | -16.9884
9.6814 | -34.3975 | -16.9981
9.6875 | -34.3979 | -17.0076
9.6936 | -34.3984 | -17.0172
9.6997 | -34.3988 | -17.0271
9.7058 | -34.3993 | -17.0368
9.7119 | -34.3997 | -17.0466
9.718 | -34.4001 | -17.0565
9.7241 | -34.4006 | -17.0662
9.7302 | -34.401 | -17.076
9.7363 | -34.4014 | -17.0857
9.7424 | -34.4019 | -17.0953
9.7485 | -34.4023 | -17.105
9.7546 | -34.4027 | -17.1146
9.7607 | -34.4031 | -17.1245
9.7668 | -34.4036 | -17.1343
9.7729 | -34.404 | -17.144
9.779 | -34.4045 | -17.1535
9.7851 | -34.4049 | -17.1634
9.7912 | -34.4054 | -17.173
9.7973 | -34.4058 | -17.1829
9.8034 | -34.4062 | -17.1924
9.8095 | -34.4067 | -17.2023
9.8157 | -34.4071 | -17.212
9.8218 | -34.4076 | -17.2219
9.8279 | -34.408 | -17.2315
9.834 | -34.4084 | -17.2413
9.8401 | -34.4089 | -17.251
9.8462 | -34.4093 | -17.2604
9.8523 | -34.4097 | -17.2702
9.8584 | -34.4102 | -17.2799
9.8645 | -34.4106 | -17.2898
9.8706 | -34.411 | -17.2996
9.8767 | -34.4114 | -17.3093
9.8828 | -34.4118 | -17.3188
9.8889 | -34.4123 | -17.3284
9.895 | -34.4127 | -17.3381
9.9011 | -34.4131 | -17.3479
9.9072 | -34.4136 | -17.3576
9.9133 | -34.414 | -17.3672
9.9194 | -34.4144 | -17.3768
9.9255 | -34.4148 | -17.3865
9.9316 | -34.4153 | -17.3962
9.9377 | -34.4157 | -17.4058
9.9438 | -34.4161 | -17.4156
9.9499 | -34.4166 | -17.4254
9.956 | -34.417 | -17.4351
9.9622 | -34.4174 | -17.4449
9.9683 | -34.4179 | -17.4546
9.9744 | -34.4183 | -17.4644
9.9805 | -34.4187 | -17.4741
9.9866 | -34.4191 | -17.4837
9.9927 | -34.4196 | -17.4933
9.9988 | -34.42 | -17.5032
10.0049 | -34.4205 | -17.5129
10.011 | -34.4209 | -17.5226
10.0171 | -34.4213 | -17.5322
10.0232 | -34.4217 | -17.5418
10.0293 | -34.4222 | -17.5515
10.0354 | -34.4226 | -17.5612
10.0415 | -34.4231 | -17.5709
10.0476 | -34.4235 | -17.5807
10.0537 | -34.4239 | -17.5905
10.0598 | -34.4244 | -17.6001
10.0659 | -34.4248 | -17.6099
10.072 | -34.4252 | -17.6194
10.0781 | -34.4256 | -17.629
10.0842 | -34.4261 | -17.6387
10.0903 | -34.4265 | -17.6484
10.0964 | -34.4269 | -17.6578
10.1026 | -34.4274 | -17.6675
10.1087 | -34.4278 | -17.6772
10.1148 | -34.4282 | -17.6869
10.1209 | -34.4286 | -17.6964
10.127 | -34.429 | -17.7059
10.1331 | -34.4295 | -17.7154
10.1392 | -34.4299 | -17.7249
10.1453 | -34.4304 | -17.7345
10.1514 | -34.4308 | -17.7442
10.1575 | -34.4312 | -17.754
10.1636 | -34.4317 | -17.7634
10.1697 | -34.4321 | -17.7732
10.1758 | -34.4325 | -17.7826
10.1819 | -34.433 | -17.7922
10.188 | -34.4334 | -17.8019
10.1941 | -34.4338 | -17.8116
10.2002 | -34.4342 | -17.8212
10.2063 | -34.4346 | -17.8308
10.2124 | -34.435 | -17.8405
10.2185 | -34.4355 | -17.8502
10.2246 | -34.4359 | -17.8597
10.2307 | -34.4363 | -17.8694
10.2368 | -34.4368 | -17.8787
10.2429 | -34.4372 | -17.8883
10.2491 | -34.4376 | -17.8981
10.2552 | -34.438 | -17.9076
10.2613 | -34.4384 | -17.9174
10.2674 | -34.4389 | -17.927
10.2735 | -34.4393 | -17.9368
10.2796 | -34.4398 | -17.9463
10.2857 | -34.4402 | -17.9561
10.2918 | -34.4406 | -17.9657
10.2979 | -34.4411 | -17.9754
10.304 | -34.4415 | -17.9849
10.3101 | -34.4419 | -17.9946
10.3162 | -34.4423 | -18.0042
10.3223 | -34.4428 | -18.0137
10.3284 | -34.4432 | -18.0233
10.3345 | -34.4436 | -18.0328
10.3406 | -34.4441 | -18.0424
10.3467 | -34.4445 | -18.052
10.3528 | -34.4449 | -18.0615
10.3589 | -34.4453 | -18.0712
10.365 | -34.4457 | -18.0809
10.3711 | -34.4462 | -18.0905
10.3772 | -34.4466 | -18.1002
10.3833 | -34.447 | -18.1097
10.3895 | -34.4475 | -18.1193
10.3956 | -34.4479 | -18.1289
10.4017 | -34.4484 | -18.1384
10.4078 | -34.4488 | -18.148
10.4139 | -34.4492 | -18.1574
10.42 | -34.4496 | -18.1671
10.4261 | -34.4501 | -18.1767
10.4322 | -34.4505 | -18.1862
10.4383 | -34.4509 | -18.1958
10.4444 | -34.4513 | -18.2052
10.4505 | -34.4517 | -18.2148
10.4566 | -34.4521 | -18.2243
10.4627 | -34.4525 | -18.2337
10.4688 | -34.453 | -18.2432
10.4749 | -34.4534 | -18.2528
10.481 | -34.4538 | -18.2624
10.4871 | -34.4543 | -18.2718
10.4932 | -34.4547 | -18.2813
10.4993 | -34.4551 | -18.2909
10.5054 | -34.4555 | -18.3005
10.5115 | -34.456 | -18.3101
10.5176 | -34.4564 | -18.3194
10.5237 | -34.4569 | -18.3292
10.5298 | -34.4573 | -18.3386
10.536 | -34.4577 | -18.3482
10.5421 | -34.4581 | -18.3576
10.5482 | -34.4585 | -18.3672
10.5543 | -34.459 | -18.3766
10.5604 | -34.4594 | -18.386
10.5665 | -34.4598 | -18.3954
10.5726 | -34.4603 | -18.4049
10.5787 | -34.4607 | -18.4143
10.5848 | -34.4611 | -18.4235
10.5909 | -34.4616 | -18.4331
10.597 | -34.462 | -18.4426
10.6031 | -34.4624 | -18.4523
10.6092 | -34.4628 | -18.4618
10.6153 | -34.4632 | -18.4714
10.6214 | -34.4637 | -18.4807
10.6275 | -34.4641 | -18.4902
10.6336 | -34.4645 | -18.4997
10.6397 | -34.4649 | -18.5092
10.6458 | -34.4653 | -18.5186
10.6519 | -34.4658 | -18.5282
10.658 | -34.4662 | -18.5376
10.6641 | -34.4666 | -18.5473
10.6702 | -34.467 | -18.5568
10.6764 | -34.4675 | -18.5663
10.6825 | -34.4679 | -18.5757
10.6886 | -34.4683 | -18.5851
10.6947 | -34.4688 | -18.5946
10.7008 | -34.4692 | -18.604
10.7069 | -34.4696 | -18.6136
10.713 | -34.4701 | -18.6231
10.7191 | -34.4705 | -18.6326
10.7252 | -34.4709 | -18.6419
10.7313 | -34.4713 | -18.6513
10.7374 | -34.4717 | -18.6607
10.7435 | -34.4722 | -18.6701
10.7496 | -34.4726 | -18.6793
10.7557 | -34.473 | -18.6889
10.7618 | -34.4734 | -18.6981
10.7679 | -34.4739 | -18.7075
10.774 | -34.4743 | -18.7171
10.7801 | -34.4747 | -18.7267
10.7862 | -34.4751 | -18.736
10.7923 | -34.4756 | -18.7455
10.7984 | -34.476 | -18.7548
10.8045 | -34.4764 | -18.764
10.8106 | -34.4769 | -18.7736
10.8168 | -34.4773 | -18.7829
10.8229 | -34.4777 | -18.7924
10.829 | -34.4781 | -18.8018
10.8351 | -34.4785 | -18.8108
10.8412 | -34.479 | -18.8201
10.8473 | -34.4794 | -18.8294
10.8534 | -34.4799 | -18.8389
10.8595 | -34.4803 | -18.8483
10.8656 | -34.4807 | -18.8575
10.8717 | -34.4811 | -18.8669
10.8778 | -34.4816 | -18.8764
10.8839 | -34.482 | -18.8858
10.89 | -34.4824 | -18.8952
10.8961 | -34.4829 | -18.9046
10.9022 | -34.4833 | -18.9139
10.9083 | -34.4837 | -18.9232
10.9144 | -34.4842 | -18.9325
10.9205 | -34.4846 | -18.9419
10.9266 | -34.485 | -18.9512
10.9327 | -34.4855 | -18.9605
10.9388 | -34.4859 | -18.9696
10.9449 | -34.4863 | -18.979
10.951 | -34.4867 | -18.9884
10.9571 | -34.4872 | -18.9978
10.9633 | -34.4876 | -19.0072
10.9694 | -34.488 | -19.0167
10.9755 | -34.4885 | -19.026
10.9816 | -34.4889 | -19.0355
10.9877 | -34.4894 | -19.0447
10.9938 | -34.4898 | -19.0542
10.9999 | -34.4902 | -19.0634
11.006 | -34.4907 | -19.0726
11.0121 | -34.4911 | -19.082
11.0182 | -34.4915 | -19.0913
11.0243 | -34.492 | -19.1007
11.0304 | -34.4924 | -19.1101
11.0365 | -34.4928 | -19.1196
11.0426 | -34.4933 | -19.1289
11.0487 | -34.4937 | -19.1384
11.0548 | -34.4942 | -19.1476
11.0609 | -34.4946 | -19.1569
11.067 | -34.495 | -19.1662
11.0731 | -34.4955 | -19.1754
11.0792 | -34.496 | -19.1846
11.0853 | -34.4964 | -19.194
11.0914 | -34.4969 | -19.2033
11.0975 | -34.4973 | -19.2126
11.1037 | -34.4977 | -19.2219
11.1098 | -34.4981 | -19.2314
11.1159 | -34.4986 | -19.2405
11.122 | -34.499 | -19.2496
11.1281 | -34.4995 | -19.259
11.1342 | -34.4999 | -19.2684
11.1403 | -34.5003 | -19.2778
11.1464 | -34.5008 | -19.2871
11.1525 | -34.5012 | -19.2961
11.1586 | -34.5016 | -19.3056
11.1647 | -34.5021 | -19.3148
11.1708 | -34.5025 | -19.3241
11.1769 | -34.5029 | -19.3334
11.183 | -34.5034 | -19.3425
11.1891 | -34.5038 | -19.3519
11.1952 | -34.5043 | -19.3611
11.2013 | -34.5048 | -19.3705
11.2074 | -34.5052 | -19.3798
11.2135 | -34.5056 | -19.3892
11.2196 | -34.5061 | -19.3983
11.2257 | -34.5065 | -19.4076
11.2318 | -34.507 | -19.4168
11.2379 | -34.5074 | -19.4261
11.244 | -34.5079 | -19.4355
11.2502 | -34.5083 | -19.4447
11.2563 | -34.5088 | -19.4538
11.2624 | -34.5092 | -19.463
11.2685 | -34.5097 | -19.4723
11.2746 | -34.5101 | -19.4814
11.2807 | -34.5105 | -19.4908
11.2868 | -34.511 | -19.5002
11.2929 | -34.5115 | -19.5092
11.299 | -34.512 | -19.5185
11.3051 | -34.5124 | -19.528
11.3112 | -34.5129 | -19.5373
11.3173 | -34.5133 | -19.5465
11.3234 | -34.5138 | -19.5555
11.3295 | -34.5142 | -19.565
11.3356 | -34.5147 | -19.5744
11.3417 | -34.5151 | -19.5835
11.3478 | -34.5155 | -19.593
11.3539 | -34.5159 | -19.6024
11.36 | -34.5164 | -19.6115
11.3661 | -34.5168 | -19.6206
11.3722 | -34.5173 | -19.6298
11.3783 | -34.5177 | -19.6392
11.3844 | -34.5182 | -19.6484
11.3906 | -34.5186 | -19.6577
11.3967 | -34.5191 | -19.6668
11.4028 | -34.5196 | -19.6762
11.4089 | -34.52 | -19.6853
11.415 | -34.5204 | -19.6946
11.4211 | -34.5209 | -19.704
11.4272 | -34.5213 | -19.7132
11.4333 | -34.5218 | -19.7221
11.4394 | -34.5222 | -19.7314
11.4455 | -34.5226 | -19.7405
11.4516 | -34.5231 | -19.7497
11.4577 | -34.5235 | -19.7588
11.4638 | -34.524 | -19.7677
11.4699 | -34.5245 | -19.777
11.476 | -34.5249 | -19.7863
11.4821 | -34.5254 | -19.7956
11.4882 | -34.5258 | -19.8049
11.4943 | -34.5262 | -19.8142
11.5004 | -34.5267 | -19.8231
11.5065 | -34.5272 | -19.8322
11.5126 | -34.5276 | -19.8412
11.5187 | -34.528 | -19.8505
11.5248 | -34.5285 | -19.8599
11.5309 | -34.529 | -19.869
11.5371 | -34.5294 | -19.878
11.5432 | -34.5299 | -19.8873
11.5493 | -34.5303 | -19.8967
11.5554 | -34.5308 | -19.906
11.5615 | -34.5312 | -19.9154
11.5676 | -34.5317 | -19.9246
11.5737 | -34.5322 | -19.9337
11.5798 | -34.5326 | -19.9426
11.5859 | -34.5331 | -19.9518
11.592 | -34.5335 | -19.9608
11.5981 | -34.534 | -19.9699
11.6042 | -34.5345 | -19.9789
11.6103 | -34.5349 | -19.9879
11.6164 | -34.5354 | -19.9971
11.6225 | -34.5358 | -20.0061
11.6286 | -34.5363 | -20.0151
11.6347 | -34.5367 | -20.0243
11.6408 | -34.5371 | -20.0336
11.6469 | -34.5376 | -20.0429
11.653 | -34.5381 | -20.052
11.6591 | -34.5385 | -20.0612
11.6652 | -34.5389 | -20.0703
11.6713 | -34.5394 | -20.0794
11.6775 | -34.5399 | -20.0884
11.6836 | -34.5403 | -20.0975
11.6897 | -34.5408 | -20.1067
11.6958 | -34.5413 | -20.1158
11.7019 | -34.5417 | -20.1252
11.708 | -34.5422 | -20.1343
11.7141 | -34.5427 | -20.1435
11.7202 | -34.5431 | -20.1527
11.7263 | -34.5436 | -20.1619
11.7324 | -34.544 | -20.1712
11.7385 | -34.5445 | -20.1804
11.7446 | -34.5449 | -20.1894
11.7507 | -34.5454 | -20.1984
11.7568 | -34.5459 | -20.2077
11.7629 | -34.5463 | -20.217
11.769 | -34.5467 | -20.2262
11.7751 | -34.5472 | -20.2353
11.7812 | -34.5477 | -20.2443
11.7873 | -34.5481 | -20.2534
11.7934 | -34.5486 | -20.2626
11.7995 | -34.5491 | -20.2718
11.8056 | -34.5495 | -20.2808
11.8117 | -34.55 | -20.29
11.8178 | -34.5505 | -20.2987
11.824 | -34.551 | -20.308
11.8301 | -34.5514 | -20.3166
11.8362 | -34.5518 | -20.326
11.8423 | -34.5523 | -20.3351
11.8484 | -34.5528 | -20.3443
11.8545 | -34.5533 | -20.3533
11.8606 | -34.5537 | -20.3624
11.8667 | -34.5542 | -20.3716
11.8728 | -34.5546 | -20.3805
11.8789 | -34.5551 | -20.3896
11.885 | -34.5556 | -20.3986
11.8911 | -34.5561 | -20.4075
11.8972 | -34.5565 | -20.4166
11.9033 | -34.557 | -20.4256
11.9094 | -34.5574 | -20.4346
11.9155 | -34.5578 | -20.4439
11.9216 | -34.5583 | -20.4528
11.9277 | -34.5588 | -20.4619
11.9338 | -34.5592 | -20.4711
11.9399 | -34.5597 | -20.4803
11.946 | -34.5602 | -20.4892
11.9521 | -34.5606 | -20.4981
11.9582 | -34.5611 | -20.5072
11.9644 | -34.5616 | -20.5162
11.9705 | -34.5621 | -20.5251
11.9766 | -34.5626 | -20.5342
11.9827 | -34.563 | -20.5433
11.9888 | -34.5635 | -20.5522
11.9949 | -34.564 | -20.561
12.001 | -34.5645 | -20.5701
12.0071 | -34.565 | -20.579
12.0132 | -34.5654 | -20.588
12.0193 | -34.5659 | -20.5971
12.0254 | -34.5663 | -20.6064
12.0315 | -34.5668 | -20.6152
12.0376 | -34.5673 | -20.6242
12.0437 | -34.5677 | -20.6332
12.0498 | -34.5682 | -20.6424
12.0559 | -34.5686 | -20.6513
12.062 | -34.5691 | -20.66
12.0681 | -34.5696 | -20.6689
12.0742 | -34.5701 | -20.6778
12.0803 | -34.5706 | -20.6868
12.0864 | -34.571 | -20.6959
12.0925 | -34.5715 | -20.7049
12.0986 | -34.572 | -20.7139
12.1047 | -34.5725 | -20.7229
12.1109 | -34.5729 | -20.7318
12.117 | -34.5734 | -20.7407
12.1231 | -34.5739 | -20.7497
12.1292 | -34.5743 | -20.7589
12.1353 | -34.5748 | -20.7678
12.1414 | -34.5753 | -20.7768
12.1475 | -34.5758 | -20.786
12.1536 | -34.5763 | -20.795
12.1597 | -34.5767 | -20.8041
12.1658 | -34.5772 | -20.8131
12.1719 | -34.5777 | -20.8219
12.178 | -34.5781 | -20.8309
12.1841 | -34.5786 | -20.84
12.1902 | -34.5791 | -20.8489
12.1963 | -34.5796 | -20.858
12.2024 | -34.58 | -20.8668
12.2085 | -34.5805 | -20.8759
12.2146 | -34.581 | -20.8847
12.2207 | -34.5814 | -20.8937
12.2268 | -34.5819 | -20.9026
12.2329 | -34.5824 | -20.9115
12.239 | -34.5829 | -20.92
12.2451 | -34.5833 | -20.9291
12.2513 | -34.5838 | -20.9383
12.2574 | -34.5843 | -20.9473
12.2635 | -34.5847 | -20.9562
12.2696 | -34.5852 | -20.9652
12.2757 | -34.5857 | -20.974
12.2818 | -34.5862 | -20.983
12.2879 | -34.5866 | -20.992
12.294 | -34.5871 | -21.0008
12.3001 | -34.5876 | -21.0098
12.3062 | -34.5881 | -21.0187
12.3123 | -34.5885 | -21.0274
12.3184 | -34.589 | -21.0362
12.3245 | -34.5895 | -21.045
12.3306 | -34.5899 | -21.0539
12.3367 | -34.5904 | -21.063
12.3428 | -34.5909 | -21.072
12.3489 | -34.5914 | -21.0805
12.355 | -34.5919 | -21.0896
12.3611 | -34.5924 | -21.0987
12.3672 | -34.5928 | -21.1076
12.3733 | -34.5933 | -21.1167
12.3794 | -34.5938 | -21.1255
12.3855 | -34.5942 | -21.1343
12.3916 | -34.5947 | -21.1432
12.3978 | -34.5952 | -21.1521
12.4039 | -34.5957 | -21.1612
12.41 | -34.5961 | -21.1702
12.4161 | -34.5966 | -21.1787
12.4222 | -34.5971 | -21.1876
12.4283 | -34.5975 | -21.1964
12.4344 | -34.598 | -21.2054
12.4405 | -34.5985 | -21.2143
12.4466 | -34.5989 | -21.2231
12.4527 | -34.5994 | -21.232
12.4588 | -34.5999 | -21.241
12.4649 | -34.6004 | -21.2498
12.471 | -34.6009 | -21.2587
12.4771 | -34.6013 | -21.2675
12.4832 | -34.6018 | -21.2765
12.4893 | -34.6023 | -21.2853
12.4954 | -34.6028 | -21.2942
12.5015 | -34.6032 | -21.303
12.5076 | -34.6037 | -21.312
12.5137 | -34.6042 | -21.3208
12.5198 | -34.6047 | -21.3295
12.5259 | -34.6051 | -21.3385
12.532 | -34.6055 | -21.3474
12.5382 | -34.606 | -21.3563
12.5443 | -34.6065 | -21.3652
12.5504 | -34.607 | -21.3741
12.5565 | -34.6074 | -21.3828
12.5626 | -34.6079 | -21.392
12.5687 | -34.6084 | -21.4009
12.5748 | -34.6089 | -21.4099
12.5809 | -34.6093 | -21.4186
12.587 | -34.6098 | -21.4276
12.5931 | -34.6103 | -21.4361
12.5992 | -34.6108 | -21.445
12.6053 | -34.6113 | -21.4537
12.6114 | -34.6117 | -21.4623
12.6175 | -34.6122 | -21.4714
12.6236 | -34.6127 | -21.4804
12.6297 | -34.6132 | -21.4891
12.6358 | -34.6137 | -21.4978
12.6419 | -34.6142 | -21.5068
12.648 | -34.6146 | -21.5156
12.6541 | -34.6151 | -21.5245
12.6602 | -34.6156 | -21.5331
12.6663 | -34.6161 | -21.542
12.6724 | -34.6166 | -21.5507
12.6785 | -34.617 | -21.5595
12.6847 | -34.6175 | -21.5681
12.6908 | -34.618 | -21.5771
12.6969 | -34.6185 | -21.586
12.703 | -34.6189 | -21.5949
12.7091 | -34.6194 | -21.6036
12.7152 | -34.6198 | -21.6123
12.7213 | -34.6203 | -21.6212
12.7274 | -34.6208 | -21.63
12.7335 | -34.6213 | -21.6385
12.7396 | -34.6218 | -21.6473
12.7457 | -34.6222 | -21.6561
12.7518 | -34.6227 | -21.6649
12.7579 | -34.6232 | -21.6737
12.764 | -34.6236 | -21.6828
12.7701 | -34.6241 | -21.6914
12.7762 | -34.6246 | -21.7001
12.7823 | -34.6251 | -21.7089
12.7884 | -34.6256 | -21.7178
12.7945 | -34.626 | -21.7268
12.8006 | -34.6265 | -21.7354
12.8067 | -34.627 | -21.7441
12.8128 | -34.6275 | -21.7528
12.8189 | -34.628 | -21.7615
12.8251 | -34.6284 | -21.7705
12.8312 | -34.6289 | -21.7795
12.8373 | -34.6294 | -21.7882
12.8434 | -34.6299 | -21.7972
12.8495 | -34.6303 | -21.8059
12.8556 | -34.6308 | -21.8146
12.8617 | -34.6313 | -21.8235
12.8678 | -34.6318 | -21.8323
12.8739 | -34.6322 | -21.8411
12.88 | -34.6327 | -21.85
12.8861 | -34.6331 | -21.8588
12.8922 | -34.6336 | -21.8676
12.8983 | -34.6341 | -21.8763
12.9044 | -34.6345 | -21.8849
12.9105 | -34.635 | -21.8936
12.9166 | -34.6355 | -21.9023
12.9227 | -34.636 | -21.9111
12.9288 | -34.6365 | -21.9199
12.9349 | -34.6369 | -21.9287
12.941 | -34.6375 | -21.9373
12.9471 | -34.6379 | -21.9461
12.9532 | -34.6384 | -21.9549
12.9593 | -34.6389 | -21.9639
12.9654 | -34.6394 | -21.9728
12.9716 | -34.6398 | -21.9816
12.9777 | -34.6403 | -21.9903
12.9838 | -34.6408 | -21.9989
12.9899 | -34.6412 | -22.0077
12.996 | -34.6417 | -22.0164
13.0021 | -34.6422 | -22.0252
13.0082 | -34.6426 | -22.0339
13.0143 | -34.6431 | -22.0424
13.0204 | -34.6436 | -22.051
13.0265 | -34.6441 | -22.0598
13.0326 | -34.6445 | -22.0686
13.0387 | -34.645 | -22.0773
13.0448 | -34.6455 | -22.0859
13.0509 | -34.6459 | -22.0946
13.057 | -34.6464 | -22.1033
13.0631 | -34.6469 | -22.1119
13.0692 | -34.6474 | -22.1205
13.0753 | -34.6479 | -22.1294
13.0814 | -34.6483 | -22.1383
13.0875 | -34.6488 | -22.1469
13.0936 | -34.6493 | -22.1556
13.0997 | -34.6498 | -22.1641
13.1058 | -34.6503 | -22.1727
13.112 | -34.6507 | -22.1814
13.1181 | -34.6512 | -22.19
13.1242 | -34.6517 | -22.1987
13.1303 | -34.6522 | -22.2075
13.1364 | -34.6526 | -22.2163
13.1425 | -34.6531 | -22.225
13.1486 | -34.6535 | -22.2336
13.1547 | -34.654 | -22.2424
13.1608 | -34.6545 | -22.251
13.1669 | -34.6549 | -22.2595
13.173 | -34.6554 | -22.2684
13.1791 | -34.6559 | -22.277
13.1852 | -34.6563 | -22.2857
13.1913 | -34.6568 | -22.2942
13.1974 | -34.6572 | -22.3029
13.2035 | -34.6577 | -22.3117
13.2096 | -34.6582 | -22.3204
13.2157 | -34.6587 | -22.329
13.2218 | -34.6591 | -22.3376
13.2279 | -34.6596 | -22.3462
13.234 | -34.6601 | -22.3551
13.2401 | -34.6606 | -22.3638
13.2462 | -34.661 | -22.3724
13.2524 | -34.6615 | -22.3812
13.2585 | -34.662 | -22.3898
13.2646 | -34.6625 | -22.3985
13.2707 | -34.663 | -22.4074
13.2768 | -34.6634 | -22.4161
13.2829 | -34.6639 | -22.4249
13.289 | -34.6644 | -22.4335
13.2951 | -34.6648 | -22.4421
13.3012 | -34.6653 | -22.4506
13.3073 | -34.6657 | -22.4595
13.3134 | -34.6662 | -22.4684
13.3195 | -34.6667 | -22.4768
13.3256 | -34.6671 | -22.4854
13.3317 | -34.6676 | -22.4941
13.3378 | -34.6681 | -22.5029
13.3439 | -34.6686 | -22.5115
13.35 | -34.6691 | -22.5202
13.3561 | -34.6695 | -22.5289
13.3622 | -34.67 | -22.5376
13.3683 | -34.6705 | -22.546
13.3744 | -34.6709 | -22.5546
13.3805 | -34.6714 | -22.5633
13.3866 | -34.6719 | -22.5717
13.3927 | -34.6724 | -22.5805
13.3989 | -34.6728 | -22.5891
13.405 | -34.6733 | -22.5979
13.4111 | -34.6737 | -22.6067
13.4172 | -34.6741 | -22.6151
13.4233 | -34.6746 | -22.6239
13.4294 | -34.6751 | -22.6326
13.4355 | -34.6755 | -22.641
13.4416 | -34.676 | -22.6496
13.4477 | -34.6765 | -22.6582
13.4538 | -34.677 | -22.6669
13.4599 | -34.6774 | -22.6754
13.466 | -34.6779 | -22.6841
13.4721 | -34.6783 | -22.693
13.4782 | -34.6788 | -22.7015
13.4843 | -34.6792 | -22.7105
13.4904 | -34.6797 | -22.7191
13.4965 | -34.6802 | -22.7279
13.5026 | -34.6806 | -22.7365
13.5087 | -34.6811 | -22.7451
13.5148 | -34.6816 | -22.7539
13.5209 | -34.682 | -22.7628
13.527 | -34.6825 | -22.7714
13.5331 | -34.683 | -22.78
13.5393 | -34.6834 | -22.7886
13.5454 | -34.6839 | -22.7972
13.5515 | -34.6844 | -22.806
13.5576 | -34.6848 | -22.8147
13.5637 | -34.6853 | -22.8235
13.5698 | -34.6858 | -22.8321
13.5759 | -34.6863 | -22.8404
13.582 | -34.6868 | -22.849
13.5881 | -34.6873 | -22.8576
13.5942 | -34.6877 | -22.8664
13.6003 | -34.6882 | -22.8749
13.6064 | -34.6886 | -22.8836
13.6125 | -34.6891 | -22.8923
13.6186 | -34.6896 | -22.9012
13.6247 | -34.6901 | -22.9099
13.6308 | -34.6905 | -22.9186
13.6369 | -34.691 | -22.9272
13.643 | -34.6915 | -22.936
13.6491 | -34.6919 | -22.9445
13.6552 | -34.6924 | -22.9531
13.6613 | -34.6929 | -22.9617
13.6674 | -34.6933 | -22.9703
13.6735 | -34.6938 | -22.9789
13.6796 | -34.6942 | -22.9875
13.6858 | -34.6947 | -22.996
13.6919 | -34.6952 | -23.0048
13.698 | -34.6956 | -23.0134
13.7041 | -34.6961 | -23.0222
13.7102 | -34.6966 | -23.0308
13.7163 | -34.697 | -23.0394
13.7224 | -34.6975 | -23.0481
13.7285 | -34.698 | -23.0567
13.7346 | -34.6984 | -23.0654
13.7407 | -34.6989 | -23.074
13.7468 | -34.6994 | -23.0827
13.7529 | -34.6998 | -23.0911
13.759 | -34.7003 | -23.0999
13.7651 | -34.7007 | -23.1087
13.7712 | -34.7012 | -23.1173
13.7773 | -34.7016 | -23.126
13.7834 | -34.7021 | -23.1345
13.7895 | -34.7026 | -23.1432
13.7956 | -34.703 | -23.152
13.8017 | -34.7035 | -23.1606
13.8078 | -34.704 | -23.1692
13.8139 | -34.7044 | -23.1778
13.82 | -34.7048 | -23.1863
13.8262 | -34.7053 | -23.1949
13.8323 | -34.7058 | -23.2037
13.8384 | -34.7062 | -23.2123
13.8445 | -34.7067 | -23.2208
13.8506 | -34.7072 | -23.2296
13.8567 | -34.7077 | -23.2382
13.8628 | -34.7081 | -23.2468
13.8689 | -34.7086 | -23.2555
13.875 | -34.7091 | -23.2639
13.8811 | -34.7095 | -23.2725
13.8872 | -34.71 | -23.2811
13.8933 | -34.7104 | -23.2895
13.8994 | -34.7109 | -23.2982
13.9055 | -34.7114 | -23.3069
13.9116 | -34.7118 | -23.3155
13.9177 | -34.7123 | -23.3242
13.9238 | -34.7128 | -23.3328
13.9299 | -34.7132 | -23.3415
13.936 | -34.7137 | -23.3502
13.9421 | -34.7142 | -23.3589
13.9482 | -34.7146 | -23.3675
13.9543 | -34.715 | -23.3762
13.9604 | -34.7155 | -23.3848
13.9665 | -34.716 | -23.3933
13.9727 | -34.7165 | -23.402
13.9788 | -34.717 | -23.4105
13.9849 | -34.7174 | -23.4191
13.991 | -34.7179 | -23.4278
13.9971 | -34.7184 | -23.4362
14.0032 | -34.7188 | -23.4449
14.0093 | -34.7193 | -23.4535
14.0154 | -34.7198 | -23.462
14.0215 | -34.7202 | -23.4704
14.0276 | -34.7207 | -23.479
14.0337 | -34.7212 | -23.4876
14.0398 | -34.7217 | -23.4962
14.0459 | -34.7221 | -23.5048
14.052 | -34.7226 | -23.5135
14.0581 | -34.723 | -23.5221
14.0642 | -34.7235 | -23.5306
14.0703 | -34.7239 | -23.5392
14.0764 | -34.7244 | -23.5477
14.0825 | -34.7249 | -23.5564
14.0886 | -34.7254 | -23.565
14.0947 | -34.7258 | -23.5736
14.1008 | -34.7263 | -23.5822
14.1069 | -34.7268 | -23.5908
14.1131 | -34.7272 | -23.5995
14.1192 | -34.7277 | -23.608
14.1253 | -34.7282 | -23.6166
14.1314 | -34.7286 | -23.6253
14.1375 | -34.7291 | -23.6338
14.1436 | -34.7296 | -23.6423
14.1497 | -34.7301 | -23.6509
14.1558 | -34.7305 | -23.6596
14.1619 | -34.731 | -23.6681
14.168 | -34.7314 | -23.6767
14.1741 | -34.7319 | -23.6851
14.1802 | -34.7324 | -23.6938
14.1863 | -34.7328 | -23.7024
14.1924 | -34.7333 | -23.7109
14.1985 | -34.7338 | -23.7194
14.2046 | -34.7342 | -23.7282
14.2107 | -34.7347 | -23.7364
14.2168 | -34.7352 | -23.7452
14.2229 | -34.7356 | -23.7539
14.229 | -34.7361 | -23.7625
14.2351 | -34.7366 | -23.7712
14.2412 | -34.7371 | -23.7798
14.2473 | -34.7375 | -23.7883
14.2534 | -34.738 | -23.7967
14.2596 | -34.7385 | -23.8053
14.2657 | -34.7389 | -23.8137
14.2718 | -34.7394 | -23.8225
14.2779 | -34.7399 | -23.8311
14.284 | -34.7404 | -23.8398
14.2901 | -34.7408 | -23.8484
14.2962 | -34.7413 | -23.857
14.3023 | -34.7417 | -23.8657
14.3084 | -34.7422 | -23.8742
14.3145 | -34.7427 | -23.8829
14.3206 | -34.7431 | -23.8917
14.3267 | -34.7436 | -23.9003
14.3328 | -34.7441 | -23.9089
14.3389 | -34.7446 | -23.9174
14.345 | -34.745 | -23.9261
14.3511 | -34.7455 | -23.9346
14.3572 | -34.746 | -23.9432
14.3633 | -34.7464 | -23.9519
14.3694 | -34.7469 | -23.9606
14.3755 | -34.7474 | -23.9692
14.3816 | -34.7479 | -23.9779
14.3877 | -34.7484 | -23.9864
14.3938 | -34.7488 | -23.9952
14.4 | -34.7493 | -24.0036
14.4061 | -34.7497 | -24.0122
14.4122 | -34.7502 | -24.0208
14.4183 | -34.7507 | -24.0293
14.4244 | -34.7511 | -24.0379
14.4305 | -34.7516 | -24.0465
14.4366 | -34.7521 | -24.055
14.4427 | -34.7526 | -24.0637
14.4488 | -34.753 | -24.0724
14.4549 | -34.7535 | -24.0808
14.461 | -34.754 | -24.0893
14.4671 | -34.7544 | -24.0981
14.4732 | -34.7549 | -24.1067
14.4793 | -34.7554 | -24.1153
14.4854 | -34.7559 | -24.1239
14.4915 | -34.7563 | -24.1322
14.4976 | -34.7568 | -24.1407
14.5037 | -34.7573 | -24.1491
14.5098 | -34.7578 | -24.1576
14.5159 | -34.7582 | -24.1662
14.522 | -34.7587 | -24.1749
14.5281 | -34.7592 | -24.1833
14.5342 | -34.7597 | -24.192
14.5403 | -34.7601 | -24.2004
14.5465 | -34.7606 | -24.2091
14.5526 | -34.7611 | -24.2178
14.5587 | -34.7615 | -24.2265
14.5648 | -34.762 | -24.2351
14.5709 | -34.7625 | -24.2436
14.577 | -34.763 | -24.2522
14.5831 | -34.7634 | -24.2607
14.5892 | -34.7639 | -24.2692
14.5953 | -34.7644 | -24.2777
14.6014 | -34.7649 | -24.2863
14.6075 | -34.7653 | -24.2949
14.6136 | -34.7658 | -24.3036
14.6197 | -34.7663 | -24.3124
14.6258 | -34.7667 | -24.321
14.6319 | -34.7672 | -24.3295
14.638 | -34.7676 | -24.338
14.6441 | -34.7681 | -24.3465
14.6502 | -34.7686 | -24.3551
14.6563 | -34.7691 | -24.3637
14.6624 | -34.7695 | -24.3724
14.6685 | -34.77 | -24.3808
14.6746 | -34.7705 | -24.3894
14.6807 | -34.771 | -24.3979
14.6869 | -34.7714 | -24.4064
14.693 | -34.7719 | -24.4149
14.6991 | -34.7724 | -24.4234
14.7052 | -34.7729 | -24.4318
14.7113 | -34.7734 | -24.4404
14.7174 | -34.7739 | -24.4489
14.7235 | -34.7743 | -24.4573
14.7296 | -34.7748 | -24.4658
14.7357 | -34.7753 | -24.4744
14.7418 | -34.7757 | -24.483
14.7479 | -34.7762 | -24.4914
14.754 | -34.7767 | -24.4999
14.7601 | -34.7771 | -24.5084
14.7662 | -34.7776 | -24.5169
14.7723 | -34.7781 | -24.5255
14.7784 | -34.7786 | -24.5341
14.7845 | -34.779 | -24.5426
14.7906 | -34.7795 | -24.5512
14.7967 | -34.78 | -24.5598
14.8028 | -34.7804 | -24.5682
14.8089 | -34.7809 | -24.5768
14.815 | -34.7814 | -24.5855
14.8211 | -34.7819 | -24.5938
14.8272 | -34.7824 | -24.6024
14.8334 | -34.7829 | -24.6108
14.8395 | -34.7833 | -24.6193
14.8456 | -34.7838 | -24.6277
14.8517 | -34.7843 | -24.6363
14.8578 | -34.7848 | -24.6448
14.8639 | -34.7853 | -24.6535
14.87 | -34.7858 | -24.662
14.8761 | -34.7863 | -24.6707
14.8822 | -34.7867 | -24.6792
14.8883 | -34.7872 | -24.6877
14.8944 | -34.7877 | -24.6962
14.9005 | -34.7881 | -24.7047
14.9066 | -34.7886 | -24.7132
14.9127 | -34.7891 | -24.7216
14.9188 | -34.7896 | -24.7302
14.9249 | -34.79 | -24.7387
14.931 | -34.7905 | -24.7472
14.9371 | -34.791 | -24.7557
14.9432 | -34.7915 | -24.7643
14.9493 | -34.792 | -24.7726
14.9554 | -34.7925 | -24.781
14.9615 | -34.7929 | -24.7896
14.9676 | -34.7934 | -24.7981
14.9738 | -34.7939 | -24.8065
14.9799 | -34.7944 | -24.8151
14.986 | -34.7949 | -24.8236
14.9921 | -34.7954 | -24.8321
14.9982 | -34.7958 | -24.8405
15.0043 | -34.7963 | -24.8488
15.0104 | -34.7968 | -24.8574
15.0165 | -34.7973 | -24.866
15.0226 | -34.7978 | -24.8744
15.0287 | -34.7982 | -24.8829
15.0348 | -34.7987 | -24.8915
15.0409 | -34.7992 | -24.9
15.047 | -34.7997 | -24.9086
15.0531 | -34.8002 | -24.9169
15.0592 | -34.8007 | -24.9255
15.0653 | -34.8012 | -24.934
15.0714 | -34.8016 | -24.9423
15.0775 | -34.8021 | -24.9508
15.0836 | -34.8026 | -24.9593
15.0897 | -34.8031 | -24.9676
15.0958 | -34.8036 | -24.976
15.1019 | -34.8041 | -24.9845
15.108 | -34.8046 | -24.9929
15.1141 | -34.805 | -25.0014
15.1203 | -34.8055 | -25.0099
15.1264 | -34.806 | -25.0186
15.1325 | -34.8065 | -25.0271
15.1386 | -34.807 | -25.0356
15.1447 | -34.8075 | -25.044
15.1508 | -34.808 | -25.0526
15.1569 | -34.8085 | -25.0609
15.163 | -34.809 | -25.069
15.1691 | -34.8094 | -25.0775
15.1752 | -34.8099 | -25.0861
15.1813 | -34.8104 | -25.0944
15.1874 | -34.8108 | -25.1029
15.1935 | -34.8113 | -25.1111
15.1996 | -34.8118 | -25.1195
15.2057 | -34.8123 | -25.1278
15.2118 | -34.8128 | -25.1363
15.2179 | -34.8133 | -25.1449
15.224 | -34.8137 | -25.1531
15.2301 | -34.8142 | -25.1615
15.2362 | -34.8147 | -25.1699
15.2423 | -34.8152 | -25.1784
15.2484 | -34.8158 | -25.1867
15.2545 | -34.8162 | -25.1951
15.2607 | -34.8168 | -25.2036
15.2668 | -34.8173 | -25.2116
15.2729 | -34.8178 | -25.2201
15.279 | -34.8183 | -25.2288
15.2851 | -34.8187 | -25.2371
15.2912 | -34.8192 | -25.2457
15.2973 | -34.8197 | -25.2543
15.3034 | -34.8202 | -25.2626
15.3095 | -34.8207 | -25.271
15.3156 | -34.8212 | -25.2793
15.3217 | -34.8217 | -25.2876
15.3278 | -34.8222 | -25.2961
15.3339 | -34.8227 | -25.3044
15.34 | -34.8232 | -25.3127
15.3461 | -34.8237 | -25.3209
15.3522 | -34.8242 | -25.3293
15.3583 | -34.8247 | -25.3379
15.3644 | -34.8252 | -25.3462
15.3705 | -34.8257 | -25.3544
15.3766 | -34.8262 | -25.3629
15.3827 | -34.8267 | -25.3712
15.3888 | -34.8272 | -25.3796
15.3949 | -34.8277 | -25.3881
15.401 | -34.8282 | -25.3967
15.4072 | -34.8287 | -25.4046
15.4133 | -34.8291 | -25.4129
15.4194 | -34.8296 | -25.4213
15.4255 | -34.8301 | -25.4297
15.4316 | -34.8306 | -25.4381
15.4377 | -34.8311 | -25.4465
15.4438 | -34.8317 | -25.4548
15.4499 | -34.8322 | -25.4631
15.456 | -34.8327 | -25.4715
15.4621 | -34.8332 | -25.4797
15.4682 | -34.8337 | -25.4883
15.4743 | -34.8342 | -25.4964
15.4804 | -34.8346 | -25.5047
15.4865 | -34.8351 | -25.5131
15.4926 | -34.8356 | -25.5215
15.4987 | -34.8361 | -25.53
15.5048 | -34.8366 | -25.5383
15.5109 | -34.8371 | -25.5466
15.517 | -34.8376 | -25.5548
15.5231 | -34.8381 | -25.563
15.5292 | -34.8386 | -25.5712
15.5353 | -34.8391 | -25.5796
15.5414 | -34.8396 | -25.588
15.5476 | -34.8401 | -25.596
15.5537 | -34.8406 | -25.6044
15.5598 | -34.8411 | -25.6128
15.5659 | -34.8416 | -25.6211
15.572 | -34.8421 | -25.6294
15.5781 | -34.8426 | -25.6376
15.5842 | -34.8431 | -25.6456
15.5903 | -34.8436 | -25.6538
15.5964 | -34.8441 | -25.6621
15.6025 | -34.8445 | -25.6703
15.6086 | -34.8451 | -25.6786
15.6147 | -34.8455 | -25.6871
15.6208 | -34.8461 | -25.6952
15.6269 | -34.8466 | -25.7034
15.633 | -34.8471 | -25.7118
15.6391 | -34.8476 | -25.7202
15.6452 | -34.8481 | -25.7286
15.6513 | -34.8486 | -25.7366
15.6574 | -34.8492 | -25.7451
15.6635 | -34.8497 | -25.7534
15.6696 | -34.8502 | -25.7617
15.6757 | -34.8507 | -25.77
15.6818 | -34.8512 | -25.7782
15.688 | -34.8517 | -25.7864
15.6941 | -34.8522 | -25.7946
15.7002 | -34.8527 | -25.8028
15.7063 | -34.8532 | -25.8109
15.7124 | -34.8537 | -25.8193
15.7185 | -34.8542 | -25.8275
15.7246 | -34.8547 | -25.836
15.7307 | -34.8552 | -25.8442
15.7368 | -34.8556 | -25.8524
15.7429 | -34.8561 | -25.8606
15.749 | -34.8566 | -25.8689
15.7551 | -34.8572 | -25.8771
15.7612 | -34.8577 | -25.8858
15.7673 | -34.8582 | -25.894
15.7734 | -34.8587 | -25.9025
15.7795 | -34.8592 | -25.9104
15.7856 | -34.8597 | -25.9185
15.7917 | -34.8602 | -25.9267
15.7978 | -34.8607 | -25.9349
15.8039 | -34.8613 | -25.9433
15.81 | -34.8617 | -25.9517
15.8161 | -34.8622 | -25.9598
15.8222 | -34.8627 | -25.9678
15.8283 | -34.8633 | -25.9756
15.8345 | -34.8638 | -25.9837
15.8406 | -34.8643 | -25.9919
15.8467 | -34.8648 | -26.0002
15.8528 | -34.8653 | -26.0084
15.8589 | -34.8658 | -26.0165
15.865 | -34.8664 | -26.0243
15.8711 | -34.8669 | -26.0327
15.8772 | -34.8674 | -26.0406
15.8833 | -34.8679 | -26.0489
15.8894 | -34.8684 | -26.0571
15.8955 | -34.8689 | -26.0652
15.9016 | -34.8694 | -26.0733
15.9077 | -34.8699 | -26.0814
15.9138 | -34.8704 | -26.0898
15.9199 | -34.8709 | -26.098
15.926 | -34.8714 | -26.1063
15.9321 | -34.8719 | -26.1142
15.9382 | -34.8725 | -26.1221
15.9443 | -34.873 | -26.1303
15.9504 | -34.8735 | -26.1387
15.9565 | -34.874 | -26.1467
15.9626 | -34.8745 | -26.155
15.9687 | -34.875 | -26.1628
15.9749 | -34.8755 | -26.171
15.981 | -34.876 | -26.1793
15.9871 | -34.8765 | -26.1875
15.9932 | -34.877 | -26.1957
15.9993 | -34.8775 | -26.2039
16.0054 | -34.878 | -26.2119
16.0115 | -34.8785 | -26.22
16.0176 | -34.879 | -26.2281
16.0237 | -34.8795 | -26.2364
16.0298 | -34.88 | -26.2444
16.0359 | -34.8805 | -26.2525
16.042 | -34.881 | -26.2605
16.0481 | -34.8815 | -26.2688
16.0542 | -34.882 | -26.2769
16.0603 | -34.8825 | -26.2849
16.0664 | -34.883 | -26.2928
16.0725 | -34.8836 | -26.3004
16.0786 | -34.8841 | -26.3083
16.0847 | -34.8846 | -26.3165
16.0908 | -34.8851 | -26.3247
16.0969 | -34.8856 | -26.3331
16.103 | -34.8861 | -26.3412
16.1091 | -34.8866 | -26.3491
16.1152 | -34.8871 | -26.3571
16.1214 | -34.8876 | -26.3654
16.1275 | -34.8881 | -26.3736
16.1336 | -34.8886 | -26.3816
16.1397 | -34.8891 | -26.3895
16.1458 | -34.8896 | -26.3975
16.1519 | -34.8901 | -26.4057
16.158 | -34.8906 | -26.4139
16.1641 | -34.8911 | -26.4219
16.1702 | -34.8916 | -26.4301
16.1763 | -34.8921 | -26.4383
16.1824 | -34.8926 | -26.4459
16.1885 | -34.8931 | -26.4538
16.1946 | -34.8936 | -26.4619
16.2007 | -34.8941 | -26.47
16.2068 | -34.8947 | -26.478
16.2129 | -34.8952 | -26.4861
16.219 | -34.8957 | -26.4941
16.2251 | -34.8962 | -26.5019
16.2312 | -34.8967 | -26.5102
16.2373 | -34.8972 | -26.5184
16.2434 | -34.8977 | -26.5265
16.2495 | -34.8983 | -26.5344
16.2556 | -34.8988 | -26.5427
16.2618 | -34.8993 | -26.5507
16.2679 | -34.8998 | -26.5589
16.274 | -34.9003 | -26.567
16.2801 | -34.9008 | -26.5752
16.2862 | -34.9013 | -26.5833
16.2923 | -34.9018 | -26.5916
16.2984 | -34.9023 | -26.5997
16.3045 | -34.9028 | -26.6078
16.3106 | -34.9033 | -26.616
16.3167 | -34.9038 | -26.6242
16.3228 | -34.9044 | -26.6322
16.3289 | -34.9049 | -26.6401
16.335 | -34.9054 | -26.6482
16.3411 | -34.9059 | -26.6563
16.3472 | -34.9064 | -26.6644
16.3533 | -34.9069 | -26.6725
16.3594 | -34.9074 | -26.6803
16.3655 | -34.9079 | -26.6884
16.3716 | -34.9084 | -26.6963
16.3777 | -34.909 | -26.7041
16.3838 | -34.9095 | -26.7122
16.3899 | -34.91 | -26.7202
16.396 | -34.9105 | -26.7283
16.4021 | -34.911 | -26.7361
16.4083 | -34.9115 | -26.744
16.4144 | -34.912 | -26.7521
16.4205 | -34.9125 | -26.76
16.4266 | -34.913 | -26.7679
16.4327 | -34.9135 | -26.7758
16.4388 | -34.914 | -26.7837
16.4449 | -34.9145 | -26.7916
16.451 | -34.915 | -26.7996
16.4571 | -34.9155 | -26.8079
16.4632 | -34.916 | -26.8158
16.4693 | -34.9165 | -26.8239
16.4754 | -34.917 | -26.8321
16.4815 | -34.9175 | -26.8401
16.4876 | -34.918 | -26.8481
16.4937 | -34.9184 | -26.8559
16.4998 | -34.919 | -26.8634
16.5059 | -34.9195 | -26.8714
16.512 | -34.92 | -26.8795
16.5181 | -34.9205 | -26.8875
16.5242 | -34.921 | -26.8956
16.5303 | -34.9215 | -26.9031
16.5364 | -34.922 | -26.911
16.5425 | -34.9225 | -26.9189
16.5487 | -34.923 | -26.9269
16.5548 | -34.9235 | -26.9345
16.5609 | -34.924 | -26.9424
16.567 | -34.9245 | -26.9498
16.5731 | -34.925 | -26.9579
16.5792 | -34.9255 | -26.9659
16.5853 | -34.926 | -26.9737
16.5914 | -34.9265 | -26.9817
16.5975 | -34.927 | -26.9898
16.6036 | -34.9275 | -26.9974
16.6097 | -34.928 | -27.0054
16.6158 | -34.9285 | -27.0134
16.6219 | -34.929 | -27.0214
16.628 | -34.9295 | -27.0293
16.6341 | -34.93 | -27.0372
16.6402 | -34.9305 | -27.0449
16.6463 | -34.931 | -27.0527
16.6524 | -34.9315 | -27.0604
16.6585 | -34.932 | -27.0683
16.6646 | -34.9326 | -27.0762
16.6707 | -34.9331 | -27.084
16.6768 | -34.9336 | -27.0919
16.6829 | -34.9341 | -27.0996
16.689 | -34.9346 | -27.1077
16.6952 | -34.9351 | -27.1156
16.7013 | -34.9356 | -27.1235
16.7074 | -34.9361 | -27.1312
16.7135 | -34.9366 | -27.139
16.7196 | -34.9371 | -27.1469
16.7257 | -34.9376 | -27.1547
16.7318 | -34.9381 | -27.1626
16.7379 | -34.9386 | -27.1704
16.744 | -34.9391 | -27.1783
16.7501 | -34.9396 | -27.1862
16.7562 | -34.9401 | -27.1943
16.7623 | -34.9406 | -27.202
16.7684 | -34.9411 | -27.2098
16.7745 | -34.9417 | -27.2178
16.7806 | -34.9422 | -27.2253
16.7867 | -34.9427 | -27.2331
16.7928 | -34.9432 | -27.2411
16.7989 | -34.9436 | -27.2491
16.805 | -34.9441 | -27.2568
16.8111 | -34.9446 | -27.2647
16.8172 | -34.9451 | -27.2731
16.8233 | -34.9456 | -27.2808
16.8294 | -34.9461 | -27.2887
16.8356 | -34.9466 | -27.2965
16.8417 | -34.9471 | -27.3043
16.8478 | -34.9476 | -27.3122
16.8539 | -34.9481 | -27.3199
16.86 | -34.9486 | -27.328
16.8661 | -34.9491 | -27.3357
16.8722 | -34.9496 | -27.3436
16.8783 | -34.9501 | -27.3514
16.8844 | -34.9506 | -27.3592
16.8905 | -34.9511 | -27.3671
16.8966 | -34.9516 | -27.3749
16.9027 | -34.9521 | -27.3829
16.9088 | -34.9526 | -27.3908
16.9149 | -34.9531 | -27.3989
16.921 | -34.9536 | -27.4065
16.9271 | -34.9541 | -27.4144
16.9332 | -34.9546 | -27.4221
16.9393 | -34.9551 | -27.4301
16.9454 | -34.9556 | -27.4378
16.9515 | -34.9561 | -27.4457
16.9576 | -34.9566 | -27.4531
16.9637 | -34.9571 | -27.4608
16.9698 | -34.9575 | -27.4685
16.9759 | -34.958 | -27.4764
16.9821 | -34.9585 | -27.4844
16.9882 | -34.959 | -27.4921
16.9943 | -34.9595 | -27.4998
17.0004 | -34.96 | -27.5075
17.0065 | -34.9605 | -27.5154
17.0126 | -34.961 | -27.5232
17.0187 | -34.9615 | -27.5311
17.0248 | -34.962 | -27.5389
17.0309 | -34.9625 | -27.5467
17.037 | -34.963 | -27.5545
17.0431 | -34.9635 | -27.5623
17.0492 | -34.9639 | -27.5703
17.0553 | -34.9644 | -27.578
17.0614 | -34.9649 | -27.5855
17.0675 | -34.9654 | -27.5935
17.0736 | -34.9659 | -27.6013
17.0797 | -34.9663 | -27.609
17.0858 | -34.9668 | -27.6168
17.0919 | -34.9673 | -27.6246
17.098 | -34.9678 | -27.6321
17.1041 | -34.9683 | -27.6398
17.1102 | -34.9688 | -27.6475
17.1163 | -34.9693 | -27.6553
17.1225 | -34.9698 | -27.6629
17.1286 | -34.9703 | -27.6703
17.1347 | -34.9708 | -27.6781
17.1408 | -34.9713 | -27.686
17.1469 | -34.9718 | -27.6937
17.153 | -34.9723 | -27.7017
17.1591 | -34.9728 | -27.7095
17.1652 | -34.9732 | -27.717
17.1713 | -34.9737 | -27.7248
17.1774 | -34.9742 | -27.7324
17.1835 | -34.9747 | -27.7401
17.1896 | -34.9752 | -27.7478
17.1957 | -34.9757 | -27.7555
17.2018 | -34.9762 | -27.7627
17.2079 | -34.9767 | -27.7706
17.214 | -34.9772 | -27.7783
17.2201 | -34.9777 | -27.7861
17.2262 | -34.9782 | -27.794
17.2323 | -34.9787 | -27.8016
17.2384 | -34.9792 | -27.8091
17.2445 | -34.9797 | -27.8168
17.2506 | -34.9802 | -27.8245
17.2567 | -34.9807 | -27.8323
17.2628 | -34.9812 | -27.84
17.269 | -34.9816 | -27.8476
17.2751 | -34.9821 | -27.8553
17.2812 | -34.9826 | -27.8631
17.2873 | -34.9831 | -27.8709
17.2934 | -34.9836 | -27.8786
17.2995 | -34.9841 | -27.8864
17.3056 | -34.9846 | -27.8939
17.3117 | -34.9851 | -27.9016
17.3178 | -34.9855 | -27.9091
17.3239 | -34.986 | -27.9167
17.33 | -34.9865 | -27.9243
17.3361 | -34.987 | -27.932
17.3422 | -34.9874 | -27.9392
17.3483 | -34.9879 | -27.9469
17.3544 | -34.9884 | -27.9547
17.3605 | -34.9888 | -27.9624
17.3666 | -34.9893 | -27.97
17.3727 | -34.9898 | -27.9778
17.3788 | -34.9903 | -27.9852
17.3849 | -34.9908 | -27.9928
17.391 | -34.9912 | -28.0005
17.3971 | -34.9917 | -28.0084
17.4032 | -34.9922 | -28.016
17.4094 | -34.9927 | -28.0235
17.4155 | -34.9932 | -28.031
17.4216 | -34.9937 | -28.0387
17.4277 | -34.9942 | -28.0465
17.4338 | -34.9947 | -28.0542
17.4399 | -34.9952 | -28.062
17.446 | -34.9956 | -28.0693
17.4521 | -34.9961 | -28.077
17.4582 | -34.9966 | -28.0845
17.4643 | -34.9971 | -28.092
17.4704 | -34.9976 | -28.0997
17.4765 | -34.9981 | -28.1074
17.4826 | -34.9986 | -28.1149
17.4887 | -34.9991 | -28.1225
17.4948 | -34.9995 | -28.1301
17.5009 | -35 | -28.1378
17.507 | -35.0005 | -28.1454
17.5131 | -35.001 | -28.1532
17.5192 | -35.0015 | -28.1605
17.5253 | -35.002 | -28.1682
17.5314 | -35.0024 | -28.1758
17.5375 | -35.0029 | -28.1834
17.5436 | -35.0033 | -28.1912
17.5497 | -35.0038 | -28.199
17.5559 | -35.0043 | -28.2063
17.562 | -35.0048 | -28.2138
17.5681 | -35.0053 | -28.2214
17.5742 | -35.0057 | -28.229
17.5803 | -35.0062 | -28.2364
17.5864 | -35.0067 | -28.244
17.5925 | -35.0072 | -28.2516
17.5986 | -35.0077 | -28.2592
17.6047 | -35.0082 | -28.267
17.6108 | -35.0087 | -28.2745
17.6169 | -35.0092 | -28.2822
17.623 | -35.0097 | -28.2896
17.6291 | -35.0101 | -28.2971
17.6352 | -35.0106 | -28.3047
17.6413 | -35.0111 | -28.3122
17.6474 | -35.0116 | -28.3198
17.6535 | -35.0121 | -28.3276
17.6596 | -35.0125 | -28.335
17.6657 | -35.013 | -28.3426
17.6718 | -35.0135 | -28.3503
17.6779 | -35.014 | -28.3579
17.684 | -35.0145 | -28.3654
17.6901 | -35.015 | -28.3731
17.6963 | -35.0155 | -28.3804
17.7024 | -35.016 | -28.3881
17.7085 | -35.0164 | -28.3956
17.7146 | -35.0169 | -28.4032
17.7207 | -35.0173 | -28.4107
17.7268 | -35.0178 | -28.4183
17.7329 | -35.0183 | -28.4259
17.739 | -35.0188 | -28.4333
17.7451 | -35.0193 | -28.4408
17.7512 | -35.0197 | -28.4484
17.7573 | -35.0202 | -28.456
17.7634 | -35.0207 | -28.4632
17.7695 | -35.0211 | -28.4706
17.7756 | -35.0216 | -28.4781
17.7817 | -35.0221 | -28.4856
17.7878 | -35.0225 | -28.4933
17.7939 | -35.023 | -28.5009
17.8 | -35.0235 | -28.5083
17.8061 | -35.0239 | -28.516
17.8122 | -35.0245 | -28.5237
17.8183 | -35.0249 | -28.5312
17.8244 | -35.0254 | -28.5387
17.8305 | -35.0259 | -28.5463
17.8366 | -35.0263 | -28.5536
17.8428 | -35.0268 | -28.5611
17.8489 | -35.0273 | -28.5685
17.855 | -35.0278 | -28.576
17.8611 | -35.0283 | -28.5834
17.8672 | -35.0288 | -28.5907
17.8733 | -35.0293 | -28.5982
17.8794 | -35.0297 | -28.6058
17.8855 | -35.0302 | -28.6132
17.8916 | -35.0306 | -28.6207
17.8977 | -35.0311 | -28.6286
17.9038 | -35.0316 | -28.6358
17.9099 | -35.0321 | -28.6434
17.916 | -35.0325 | -28.651
17.9221 | -35.033 | -28.6586
17.9282 | -35.0335 | -28.666
17.9343 | -35.0339 | -28.6735
17.9404 | -35.0344 | -28.6808
17.9465 | -35.0349 | -28.6884
17.9526 | -35.0354 | -28.6959
17.9587 | -35.0358 | -28.7034
17.9648 | -35.0363 | -28.7108
17.9709 | -35.0368 | -28.7183
17.977 | -35.0372 | -28.7256
17.9832 | -35.0377 | -28.7332
17.9893 | -35.0382 | -28.7407
17.9954 | -35.0387 | -28.7482
18.0015 | -35.0391 | -28.7558
18.0076 | -35.0396 | -28.763
18.0137 | -35.0401 | -28.7705
18.0198 | -35.0405 | -28.778
18.0259 | -35.041 | -28.7855
18.032 | -35.0415 | -28.7929
18.0381 | -35.0419 | -28.8003
18.0442 | -35.0424 | -28.8078
18.0503 | -35.0429 | -28.8152
18.0564 | -35.0433 | -28.8226
18.0625 | -35.0438 | -28.8303
18.0686 | -35.0443 | -28.8377
18.0747 | -35.0447 | -28.8452
18.0808 | -35.0452 | -28.8526
18.0869 | -35.0457 | -28.8602
18.093 | -35.0462 | -28.8676
18.0991 | -35.0466 | -28.8751
18.1052 | -35.0471 | -28.8825
18.1113 | -35.0476 | -28.89
18.1174 | -35.0481 | -28.8972
18.1236 | -35.0485 | -28.9049
18.1297 | -35.049 | -28.9128
18.1358 | -35.0495 | -28.9195
18.1419 | -35.05 | -28.9265
18.148 | -35.0504 | -28.9341
18.1541 | -35.0509 | -28.941
18.1602 | -35.0514 | -28.9484
18.1663 | -35.0518 | -28.9558
18.1724 | -35.0523 | -28.9632
18.1785 | -35.0527 | -28.9706
18.1846 | -35.0532 | -28.978
18.1907 | -35.0537 | -28.9853
18.1968 | -35.0541 | -28.9928
18.2029 | -35.0546 | -29.0003
18.209 | -35.055 | -29.0075
18.2151 | -35.0555 | -29.0151
18.2212 | -35.056 | -29.0225
18.2273 | -35.0565 | -29.0299
18.2334 | -35.057 | -29.0372
18.2395 | -35.0574 | -29.0446
18.2456 | -35.0579 | -29.052
18.2517 | -35.0584 | -29.0593
18.2578 | -35.0588 | -29.0665
18.2639 | -35.0593 | -29.0739
18.2701 | -35.0598 | -29.0813
18.2762 | -35.0603 | -29.0886
18.2823 | -35.0607 | -29.0961
18.2884 | -35.0612 | -29.1035
18.2945 | -35.0617 | -29.1107
18.3006 | -35.0621 | -29.118
18.3067 | -35.0626 | -29.1253
18.3128 | -35.0631 | -29.1327
18.3189 | -35.0636 | -29.1401
18.325 | -35.0641 | -29.1471
18.3311 | -35.0646 | -29.1546
18.3372 | -35.0651 | -29.1619
18.3433 | -35.0655 | -29.1692
18.3494 | -35.066 | -29.1765
18.3555 | -35.0664 | -29.184
18.3616 | -35.0669 | -29.1912
18.3677 | -35.0674 | -29.1987
18.3738 | -35.0679 | -29.206
18.3799 | -35.0683 | -29.2133
18.386 | -35.0688 | -29.2207
18.3921 | -35.0692 | -29.2281
18.3982 | -35.0697 | -29.2353
18.4043 | -35.0702 | -29.2425
18.4105 | -35.0707 | -29.25
18.4166 | -35.0711 | -29.2574
18.4227 | -35.0716 | -29.2648
18.4288 | -35.072 | -29.2723
18.4349 | -35.0725 | -29.2794
18.441 | -35.073 | -29.2866
18.4471 | -35.0734 | -29.2938
18.4532 | -35.0739 | -29.3012
18.4593 | -35.0744 | -29.3085
18.4654 | -35.0748 | -29.3151
18.4715 | -35.0753 | -29.3225
18.4776 | -35.0757 | -29.33
18.4837 | -35.0762 | -29.3372
18.4898 | -35.0767 | -29.3445
18.4959 | -35.0771 | -29.3521
18.502 | -35.0776 | -29.3592
18.5081 | -35.0781 | -29.3669
18.5142 | -35.0785 | -29.3743
18.5203 | -35.079 | -29.3815
18.5264 | -35.0795 | -29.3891
18.5325 | -35.08 | -29.3966
18.5386 | -35.0805 | -29.4042
18.5447 | -35.081 | -29.4115
18.5508 | -35.0814 | -29.4186
18.557 | -35.0819 | -29.4258
18.5631 | -35.0823 | -29.4331
18.5692 | -35.0827 | -29.4404
18.5753 | -35.0833 | -29.4475
18.5814 | -35.0837 | -29.4548
18.5875 | -35.0842 | -29.4619
18.5936 | -35.0847 | -29.4694
18.5997 | -35.0851 | -29.4768
18.6058 | -35.0856 | -29.484
18.6119 | -35.0861 | -29.4913
18.618 | -35.0866 | -29.4985
18.6241 | -35.0871 | -29.506
18.6302 | -35.0875 | -29.5133
18.6363 | -35.0879 | -29.5206
18.6424 | -35.0884 | -29.5277
18.6485 | -35.0889 | -29.5351
18.6546 | -35.0893 | -29.5422
18.6607 | -35.0898 | -29.5494
18.6668 | -35.0902 | -29.5566
18.6729 | -35.0907 | -29.5639
18.679 | -35.0912 | -29.571
18.6851 | -35.0917 | -29.578
18.6912 | -35.0921 | -29.5851
18.6974 | -35.0926 | -29.5924
18.7035 | -35.0931 | -29.5999
18.7096 | -35.0936 | -29.6073
18.7157 | -35.0941 | -29.6146
18.7218 | -35.0946 | -29.6218
18.7279 | -35.0951 | -29.629
18.734 | -35.0956 | -29.6362
18.7401 | -35.0961 | -29.6434
18.7462 | -35.0965 | -29.6504
18.7523 | -35.097 | -29.6577
18.7584 | -35.0975 | -29.6652
18.7645 | -35.098 | -29.6722
18.7706 | -35.0984 | -29.6794
18.7767 | -35.0989 | -29.6864
18.7828 | -35.0994 | -29.6934
18.7889 | -35.0998 | -29.7007
18.795 | -35.1004 | -29.7078
18.8011 | -35.1008 | -29.715
18.8072 | -35.1013 | -29.7222
18.8133 | -35.1018 | -29.7297
18.8194 | -35.1023 | -29.7365
18.8255 | -35.1028 | -29.744
18.8316 | -35.1033 | -29.7512
18.8377 | -35.1037 | -29.7585
18.8439 | -35.1042 | -29.7656
18.85 | -35.1046 | -29.7728
18.8561 | -35.1051 | -29.78
18.8622 | -35.1056 | -29.7871
18.8683 | -35.1061 | -29.794
18.8744 | -35.1066 | -29.8009
18.8805 | -35.1071 | -29.8082
18.8866 | -35.1075 | -29.8155
18.8927 | -35.1081 | -29.8225
18.8988 | -35.1086 | -29.8296
18.9049 | -35.109 | -29.8366
18.911 | -35.1095 | -29.8444
18.9171 | -35.1099 | -29.851
18.9232 | -35.1101 | -29.8577
18.9293 | -35.1107 | -29.8646
18.9354 | -35.1112 | -29.8716
18.9415 | -35.1117 | -29.8789
18.9476 | -35.1122 | -29.886
18.9537 | -35.1126 | -29.8931
18.9598 | -35.1131 | -29.9003
18.9659 | -35.1136 | -29.9076
18.972 | -35.1141 | -29.9147
18.9781 | -35.1145 | -29.922
18.9843 | -35.115 | -29.9291
18.9904 | -35.1155 | -29.9363
18.9965 | -35.116 | -29.9432
19.0026 | -35.1165 | -29.9505
19.0087 | -35.117 | -29.9577
19.0148 | -35.1174 | -29.9649
19.0209 | -35.1179 | -29.9721
19.027 | -35.1184 | -29.9792
19.0331 | -35.1189 | -29.9861
19.0392 | -35.1194 | -29.9932
19.0453 | -35.1199 | -30.0002
19.0514 | -35.1204 | -30.0071
19.0575 | -35.1209 | -30.0142
19.0636 | -35.1213 | -30.0212
19.0697 | -35.1218 | -30.0285
19.0758 | -35.1223 | -30.0355
19.0819 | -35.1227 | -30.0428
19.088 | -35.1232 | -30.05
19.0941 | -35.1237 | -30.0571
19.1002 | -35.1242 | -30.0643
19.1063 | -35.1247 | -30.0712
19.1124 | -35.1251 | -30.0783
19.1185 | -35.1256 | -30.0854
19.1246 | -35.1261 | -30.0928
19.1308 | -35.1266 | -30.0999
19.1369 | -35.1271 | -30.1071
19.143 | -35.1276 | -30.1139
19.1491 | -35.128 | -30.1212
19.1552 | -35.1285 | -30.1285
19.1613 | -35.129 | -30.1355
19.1674 | -35.1295 | -30.1424
19.1735 | -35.13 | -30.1494
19.1796 | -35.1305 | -30.1566
19.1857 | -35.131 | -30.1639
19.1918 | -35.1315 | -30.1712
19.1979 | -35.1319 | -30.1782
19.204 | -35.1325 | -30.1851
19.2101 | -35.1329 | -30.1921
19.2162 | -35.1334 | -30.199
19.2223 | -35.1339 | -30.2061
19.2284 | -35.1344 | -30.2133
19.2345 | -35.1349 | -30.2205
19.2406 | -35.1353 | -30.2272
19.2467 | -35.1358 | -30.2343
19.2528 | -35.1363 | -30.2416
19.2589 | -35.1367 | -30.2486
19.265 | -35.1372 | -30.2558
19.2712 | -35.1377 | -30.2628
19.2773 | -35.1381 | -30.2697
19.2834 | -35.1386 | -30.2767
19.2895 | -35.1391 | -30.2837
19.2956 | -35.1395 | -30.2908
19.3017 | -35.14 | -30.2977
19.3078 | -35.1405 | -30.3048
19.3139 | -35.141 | -30.3116
19.32 | -35.1415 | -30.3187
19.3261 | -35.142 | -30.3261
19.3322 | -35.1425 | -30.3332
19.3383 | -35.1429 | -30.3401
19.3444 | -35.1434 | -30.3471
19.3505 | -35.1439 | -30.354
19.3566 | -35.1444 | -30.361
19.3627 | -35.1449 | -30.3682
19.3688 | -35.1453 | -30.3752
19.3749 | -35.1458 | -30.3823
19.381 | -35.1464 | -30.389
19.3871 | -35.1469 | -30.3958
19.3932 | -35.1473 | -30.403
19.3993 | -35.1478 | -30.4103
19.4054 | -35.1483 | -30.4175
19.4115 | -35.1487 | -30.4245
19.4177 | -35.1492 | -30.4315
19.4238 | -35.1497 | -30.4384
19.4299 | -35.1502 | -30.4457
19.436 | -35.1507 | -30.453
19.4421 | -35.1512 | -30.4602
19.4482 | -35.1517 | -30.4675
19.4543 | -35.1522 | -30.4743
19.4604 | -35.1527 | -30.4811
19.4665 | -35.1531 | -30.488
19.4726 | -35.1536 | -30.4948
19.4787 | -35.1541 | -30.5016
19.4848 | -35.1545 | -30.5087
19.4909 | -35.1551 | -30.5158
19.497 | -35.1556 | -30.5228
19.5031 | -35.156 | -30.5299
19.5092 | -35.1565 | -30.5369
19.5153 | -35.157 | -30.544
19.5214 | -35.1575 | -30.5507
19.5275 | -35.158 | -30.5576
19.5336 | -35.1584 | -30.5644
19.5397 | -35.1589 | -30.5714
19.5458 | -35.1594 | -30.5785
19.5519 | -35.1599 | -30.5855
19.5581 | -35.1604 | -30.5923
19.5642 | -35.1609 | -30.5991
19.5703 | -35.1613 | -30.6063
19.5764 | -35.1619 | -30.6132
19.5825 | -35.1623 | -30.6202
19.5886 | -35.1628 | -30.6272
19.5947 | -35.1633 | -30.6341
19.6008 | -35.1638 | -30.6411
19.6069 | -35.1643 | -30.6483
19.613 | -35.1648 | -30.6553
19.6191 | -35.1652 | -30.6622
19.6252 | -35.1657 | -30.6692
19.6313 | -35.1662 | -30.6762
19.6374 | -35.1667 | -30.6831
19.6435 | -35.1672 | -30.6899
19.6496 | -35.1677 | -30.697
19.6557 | -35.1681 | -30.7039
19.6618 | -35.1687 | -30.7109
19.6679 | -35.1692 | -30.7181
19.674 | -35.1697 | -30.7251
19.6801 | -35.1702 | -30.7321
19.6862 | -35.1707 | -30.7392
19.6923 | -35.1711 | -30.7463
19.6984 | -35.1715 | -30.7531
19.7046 | -35.172 | -30.7606
19.7107 | -35.1726 | -30.7675
19.7168 | -35.173 | -30.7755
19.7229 | -35.1736 | -30.7829
19.729 | -35.1741 | -30.7902
19.7351 | -35.1746 | -30.7968
19.7412 | -35.1751 | -30.8036
19.7473 | -35.1756 | -30.8105
19.7534 | -35.1761 | -30.8176
19.7595 | -35.1766 | -30.8247
19.7656 | -35.177 | -30.8315
19.7717 | -35.1775 | -30.8388
19.7778 | -35.178 | -30.8455
19.7839 | -35.1785 | -30.8522
19.79 | -35.179 | -30.8589
19.7961 | -35.1794 | -30.8658
19.8022 | -35.1799 | -30.8726
19.8083 | -35.1804 | -30.8799
19.8144 | -35.1809 | -30.8869
19.8205 | -35.1814 | -30.8938
19.8266 | -35.1819 | -30.9006
19.8327 | -35.1824 | -30.9076
19.8388 | -35.1828 | -30.9143
19.845 | -35.1834 | -30.9213
19.8511 | -35.1839 | -30.9282
19.8572 | -35.1843 | -30.9353
19.8633 | -35.1848 | -30.9424
19.8694 | -35.1853 | -30.9495
19.8755 | -35.1857 | -30.9564
19.8816 | -35.1862 | -30.9634
19.8877 | -35.1866 | -30.9703
19.8938 | -35.1871 | -30.9772
19.8999 | -35.1876 | -30.9842
19.906 | -35.188 | -30.9909
19.9121 | -35.1885 | -30.9977
19.9182 | -35.189 | -31.0046
19.9243 | -35.1895 | -31.0119
19.9304 | -35.19 | -31.0188
19.9365 | -35.1905 | -31.0258
19.9426 | -35.191 | -31.0328
19.9487 | -35.1915 | -31.0398
19.9548 | -35.192 | -31.0467
19.9609 | -35.1925 | -31.0538
19.967 | -35.193 | -31.0609
19.9731 | -35.1935 | -31.0681
19.9792 | -35.194 | -31.0748
19.9853 | -35.1945 | -31.0815
19.9915 | -35.1949 | -31.0884
19.9976 | -35.1954 | -31.0952
20.0037 | -35.1959 | -31.1021
20.0098 | -35.1964 | -31.1089
20.0159 | -35.197 | -31.1158
20.022 | -35.1975 | -31.1227
20.0281 | -35.198 | -31.1297
20.0342 | -35.1984 | -31.1365
20.0403 | -35.199 | -31.1433
20.0464 | -35.1995 | -31.1501
20.0525 | -35.2 | -31.1571
20.0586 | -35.2005 | -31.1639
20.0647 | -35.201 | -31.1707
20.0708 | -35.2015 | -31.1776
20.0769 | -35.202 | -31.1846
20.083 | -35.2025 | -31.1915
20.0891 | -35.203 | -31.1984
20.0952 | -35.2035 | -31.2052
20.1013 | -35.204 | -31.2121
20.1074 | -35.2045 | -31.2189
20.1135 | -35.2049 | -31.2259
20.1196 | -35.2054 | -31.2326
20.1257 | -35.2059 | -31.2395
20.1319 | -35.2064 | -31.2464
20.138 | -35.2069 | -31.2533
20.1441 | -35.2074 | -31.2603
20.1502 | -35.2079 | -31.2672
20.1563 | -35.2083 | -31.2742
20.1624 | -35.2088 | -31.2812
20.1685 | -35.2093 | -31.2881
20.1746 | -35.2098 | -31.2951
20.1807 | -35.2103 | -31.3023
20.1868 | -35.2108 | -31.3094
20.1929 | -35.2112 | -31.3165
20.199 | -35.2116 | -31.3238
20.2051 | -35.2121 | -31.3306
20.2112 | -35.2126 | -31.3375
20.2173 | -35.2131 | -31.3444
20.2234 | -35.2136 | -31.3513
20.2295 | -35.2141 | -31.358
20.2356 | -35.2147 | -31.3651
20.2417 | -35.2151 | -31.372
20.2478 | -35.2156 | -31.3789
20.2539 | -35.2161 | -31.386
20.26 | -35.2166 | -31.3925
20.2661 | -35.2171 | -31.3995
20.2723 | -35.2176 | -31.4064
20.2784 | -35.2181 | -31.4133
20.2845 | -35.2186 | -31.4201
20.2906 | -35.2191 | -31.4271
20.2967 | -35.2196 | -31.4341
20.3028 | -35.2201 | -31.4407
20.3089 | -35.2206 | -31.448
20.315 | -35.221 | -31.4548
20.3211 | -35.2215 | -31.4616
20.3272 | -35.222 | -31.4685
20.3333 | -35.2224 | -31.4749
20.3394 | -35.223 | -31.4818
20.3455 | -35.2235 | -31.4888
20.3516 | -35.224 | -31.4957
20.3577 | -35.2245 | -31.5027
20.3638 | -35.225 | -31.5094
20.3699 | -35.2255 | -31.5158
20.376 | -35.226 | -31.5226
20.3821 | -35.2265 | -31.5293
20.3882 | -35.227 | -31.5363
20.3943 | -35.2275 | -31.5431
20.4004 | -35.228 | -31.55
20.4065 | -35.2285 | -31.5571
20.4126 | -35.229 | -31.5642
20.4188 | -35.2294 | -31.5709
20.4249 | -35.2299 | -31.5778
20.431 | -35.2304 | -31.5846
20.4371 | -35.231 | -31.5913
20.4432 | -35.2314 | -31.5984
20.4493 | -35.2319 | -31.6054
20.4554 | -35.2324 | -31.6122
20.4615 | -35.2329 | -31.6196
20.4676 | -35.2334 | -31.6266
20.4737 | -35.2339 | -31.6334
20.4798 | -35.2344 | -31.64
20.4859 | -35.2349 | -31.6471
20.492 | -35.2354 | -31.654
20.4981 | -35.2359 | -31.6611
20.5042 | -35.2364 | -31.6681
20.5103 | -35.2369 | -31.6748
20.5164 | -35.2373 | -31.6811
20.5225 | -35.2378 | -31.688
20.5286 | -35.2383 | -31.6948
20.5347 | -35.2388 | -31.7019
20.5408 | -35.2393 | -31.7087
20.5469 | -35.2398 | -31.716
20.553 | -35.2404 | -31.7228
20.5592 | -35.2409 | -31.7297
20.5653 | -35.2413 | -31.7365
20.5714 | -35.2418 | -31.7434
20.5775 | -35.2423 | -31.75
20.5836 | -35.2428 | -31.7567
20.5897 | -35.2433 | -31.7635
20.5958 | -35.2438 | -31.7707
20.6019 | -35.2443 | -31.7778
20.608 | -35.2449 | -31.7846
20.6141 | -35.2454 | -31.7914
20.6202 | -35.2459 | -31.7985
20.6263 | -35.2464 | -31.8052
20.6324 | -35.2469 | -31.812
20.6385 | -35.2473 | -31.819
20.6446 | -35.2478 | -31.8258
20.6507 | -35.2483 | -31.8327
20.6568 | -35.2488 | -31.8396
20.6629 | -35.2492 | -31.8466
20.669 | -35.2497 | -31.8534
20.6751 | -35.2501 | -31.8601
20.6812 | -35.2507 | -31.8666
20.6873 | -35.2512 | -31.873
20.6934 | -35.2516 | -31.8799
20.6995 | -35.2522 | -31.8868
20.7057 | -35.2526 | -31.8936
20.7118 | -35.2531 | -31.9004
20.7179 | -35.2537 | -31.9068
20.724 | -35.2541 | -31.9138
20.7301 | -35.2546 | -31.9208
20.7362 | -35.2551 | -31.9277
20.7423 | -35.2555 | -31.9346
20.7484 | -35.256 | -31.9416
20.7545 | -35.2565 | -31.9485
20.7606 | -35.2571 | -31.9552
20.7667 | -35.2576 | -31.9622
20.7728 | -35.258 | -31.9688
20.7789 | -35.2586 | -31.9758
20.785 | -35.2591 | -31.9829
20.7911 | -35.2596 | -31.9892
20.7972 | -35.2601 | -31.9959
20.8033 | -35.2606 | -32.0027
20.8094 | -35.261 | -32.0097
20.8155 | -35.2615 | -32.0161
20.8216 | -35.262 | -32.023
20.8277 | -35.2625 | -32.0296
20.8338 | -35.263 | -32.0364
20.8399 | -35.2635 | -32.0432
20.8461 | -35.264 | -32.0497
20.8522 | -35.2645 | -32.0564
20.8583 | -35.2649 | -32.0634
20.8644 | -35.2654 | -32.0701
20.8705 | -35.2658 | -32.0771
20.8766 | -35.2663 | -32.0839
20.8827 | -35.2668 | -32.0907
20.8888 | -35.2673 | -32.0973
20.8949 | -35.2678 | -32.1036
20.901 | -35.2682 | -32.1107
20.9071 | -35.2688 | -32.1177
20.9132 | -35.2692 | -32.1246
20.9193 | -35.2697 | -32.1315
20.9254 | -35.2702 | -32.1383
20.9315 | -35.2707 | -32.145
20.9376 | -35.2712 | -32.152
20.9437 | -35.2717 | -32.1586
20.9498 | -35.2722 | -32.1654
20.9559 | -35.2727 | -32.1719
20.962 | -35.2732 | -32.1788
20.9681 | -35.2737 | -32.1853
20.9742 | -35.2742 | -32.1923
20.9803 | -35.2747 | -32.1989
20.9864 | -35.2751 | -32.2061
20.9926 | -35.2755 | -32.2129
20.9987 | -35.276 | -32.2194
21.0048 | -35.2765 | -32.2267
21.0109 | -35.277 | -32.2332
21.017 | -35.2776 | -32.2401
21.0231 | -35.278 | -32.247
21.0292 | -35.2785 | -32.2537
21.0353 | -35.2791 | -32.2598
21.0414 | -35.2795 | -32.2664
21.0475 | -35.28 | -32.2734
21.0536 | -35.2805 | -32.2802
21.0597 | -35.2809 | -32.2874
21.0658 | -35.2814 | -32.2943
21.0719 | -35.2819 | -32.3016
21.078 | -35.2825 | -32.3085
21.0841 | -35.283 | -32.315
21.0902 | -35.2835 | -32.3216
21.0963 | -35.2839 | -32.3281
21.1024 | -35.2844 | -32.3351
21.1085 | -35.2848 | -32.3422
21.1146 | -35.2854 | -32.3491
21.1207 | -35.2859 | -32.3561
21.1268 | -35.2864 | -32.3632
21.133 | -35.2869 | -32.37
21.1391 | -35.2874 | -32.3767
21.1452 | -35.288 | -32.3836
21.1513 | -35.2885 | -32.3906
21.1574 | -35.289 | -32.3971
21.1635 | -35.2895 | -32.4036
21.1696 | -35.29 | -32.4105
21.1757 | -35.2905 | -32.4174
21.1818 | -35.2911 | -32.4242
21.1879 | -35.2916 | -32.4309
21.194 | -35.2921 | -32.4378
21.2001 | -35.2926 | -32.4444
21.2062 | -35.2932 | -32.451
21.2123 | -35.2935 | -32.4572
21.2184 | -35.294 | -32.4639
21.2245 | -35.2945 | -32.4708
21.2306 | -35.295 | -32.4773
21.2367 | -35.2955 | -32.4846
21.2428 | -35.2959 | -32.4914
21.2489 | -35.2963 | -32.4984
21.255 | -35.2967 | -32.5053
21.2611 | -35.2973 | -32.5121
21.2672 | -35.2978 | -32.5192
21.2733 | -35.2984 | -32.526
21.2795 | -35.2986 | -32.5329
21.2856 | -35.2989 | -32.5389
21.2917 | -35.2993 | -32.5467
21.2978 | -35.2996 | -32.5533
21.3039 | -35.3 | -32.5597
21.31 | -35.3005 | -32.5671
21.3161 | -35.301 | -32.573
21.3222 | -35.3014 | -32.5803
21.3283 | -35.3019 | -32.5867
21.3344 | -35.3025 | -32.5932
21.3405 | -35.303 | -32.6
21.3466 | -35.3035 | -32.607
21.3527 | -35.304 | -32.6139
21.3588 | -35.3044 | -32.6207
21.3649 | -35.3049 | -32.6276
21.371 | -35.3054 | -32.6345
21.3771 | -35.3059 | -32.6414
21.3832 | -35.3064 | -32.6481
21.3893 | -35.3068 | -32.6543
21.3954 | -35.3072 | -32.6611
21.4015 | -35.3077 | -32.6677
21.4076 | -35.3081 | -32.6748
21.4137 | -35.3086 | -32.6817
21.4199 | -35.3091 | -32.6887
21.426 | -35.3097 | -32.696
21.4321 | -35.3101 | -32.7023
21.4382 | -35.3106 | -32.7091
21.4443 | -35.3111 | -32.7162
21.4504 | -35.3116 | -32.7235
21.4565 | -35.312 | -32.7301
21.4626 | -35.3124 | -32.7366
21.4687 | -35.3129 | -32.7434
21.4748 | -35.3133 | -32.7495
21.4809 | -35.3138 | -32.7565
21.487 | -35.3142 | -32.7634
21.4931 | -35.3148 | -32.7695
21.4992 | -35.3152 | -32.7759
21.5053 | -35.3157 | -32.7827
21.5114 | -35.316 | -32.7896
21.5175 | -35.3164 | -32.7971
21.5236 | -35.3168 | -32.8033
21.5297 | -35.3174 | -32.8102
21.5358 | -35.3178 | -32.8171
21.5419 | -35.3183 | -32.8241
21.548 | -35.3188 | -32.8306
21.5541 | -35.3193 | -32.8376
21.5602 | -35.3198 | -32.8448
21.5664 | -35.3202 | -32.8518
21.5725 | -35.3206 | -32.8585
21.5786 | -35.3211 | -32.8657
21.5847 | -35.3214 | -32.8727
21.5908 | -35.3219 | -32.8796
21.5969 | -35.3223 | -32.8862
21.603 | -35.3228 | -32.8927
21.6091 | -35.3233 | -32.8997
21.6152 | -35.3237 | -32.9065
21.6213 | -35.3241 | -32.9138
21.6274 | -35.3245 | -32.9209
21.6335 | -35.325 | -32.9269
21.6396 | -35.3254 | -32.9337
21.6457 | -35.3259 | -32.9412
21.6518 | -35.3264 | -32.9483
21.6579 | -35.3268 | -32.955
21.664 | -35.3272 | -32.9614
21.6701 | -35.3276 | -32.9683
21.6762 | -35.3281 | -32.9752
21.6823 | -35.3286 | -32.982
21.6884 | -35.329 | -32.9888
21.6945 | -35.3295 | -32.9951
21.7006 | -35.33 | -33.0018
21.7068 | -35.3304 | -33.0089
21.7129 | -35.3309 | -33.0151
21.719 | -35.3313 | -33.0227
21.7251 | -35.3318 | -33.0298
21.7312 | -35.3322 | -33.0361
21.7373 | -35.3328 | -33.0427
21.7434 | -35.3332 | -33.0498
21.7495 | -35.3336 | -33.0569
21.7556 | -35.334 | -33.0636
21.7617 | -35.3344 | -33.0699
21.7678 | -35.3349 | -33.0769
21.7739 | -35.3353 | -33.0835
21.78 | -35.3358 | -33.0905
21.7861 | -35.3363 | -33.0974
21.7922 | -35.3367 | -33.1034
21.7983 | -35.3371 | -33.1105
21.8044 | -35.3376 | -33.1171
21.8105 | -35.3379 | -33.1237
21.8166 | -35.3383 | -33.1308
21.8227 | -35.3387 | -33.1378
21.8288 | -35.3391 | -33.144
21.8349 | -35.3396 | -33.1507
21.841 | -35.3399 | -33.1569
21.8471 | -35.3404 | -33.163
21.8533 | -35.3409 | -33.1696
21.8594 | -35.3413 | -33.1768
21.8655 | -35.3417 | -33.1829
21.8716 | -35.3423 | -33.1892
21.8777 | -35.3428 | -33.1965
21.8838 | -35.3433 | -33.2036
21.8899 | -35.3437 | -33.2106
21.896 | -35.3443 | -33.2173
21.9021 | -35.3447 | -33.2238
21.9082 | -35.3452 | -33.2314
21.9143 | -35.3455 | -33.2385
21.9204 | -35.346 | -33.245
21.9265 | -35.3464 | -33.2518
21.9326 | -35.3469 | -33.2591
21.9387 | -35.3474 | -33.2659
21.9448 | -35.3477 | -33.2732
21.9509 | -35.3481 | -33.2793
21.957 | -35.3485 | -33.2867
21.9631 | -35.3489 | -33.2941
21.9692 | -35.3494 | -33.3008
21.9753 | -35.3498 | -33.3077
21.9814 | -35.3502 | -33.3147
21.9875 | -35.3506 | -33.3208
21.9937 | -35.351 | -33.3276
21.9998 | -35.3514 | -33.3338
22.0059 | -35.3519 | -33.3409
22.012 | -35.3524 | -33.3476
22.0181 | -35.3529 | -33.3535
22.0242 | -35.3533 | -33.3596
22.0303 | -35.3537 | -33.3658
22.0364 | -35.354 | -33.3728
22.0425 | -35.3545 | -33.3792
22.0486 | -35.3549 | -33.3855
22.0547 | -35.3553 | -33.3927
22.0608 | -35.3557 | -33.3995
22.0669 | -35.3562 | -33.4062
22.073 | -35.3567 | -33.4122
22.0791 | -35.3572 | -33.4186
22.0852 | -35.3577 | -33.4255
22.0913 | -35.3581 | -33.4317
22.0974 | -35.3585 | -33.4387
22.1035 | -35.3589 | -33.4453
22.1096 | -35.3593 | -33.4517
22.1157 | -35.3597 | -33.4581
22.1218 | -35.3602 | -33.4648
22.1279 | -35.3605 | -33.4705
22.134 | -35.361 | -33.4773
22.1402 | -35.3615 | -33.4839
22.1463 | -35.3619 | -33.4905
22.1524 | -35.3624 | -33.4973
22.1585 | -35.3628 | -33.5039
22.1646 | -35.3632 | -33.5104
22.1707 | -35.3637 | -33.518
22.1768 | -35.3641 | -33.5247
22.1829 | -35.3645 | -33.531
22.189 | -35.365 | -33.5378
22.1951 | -35.3653 | -33.5444
22.2012 | -35.3657 | -33.5513
22.2073 | -35.3662 | -33.5579
22.2134 | -35.3666 | -33.564
22.2195 | -35.3671 | -33.5708
22.2256 | -35.3676 | -33.5774
22.2317 | -35.368 | -33.5845
22.2378 | -35.3684 | -33.5914
22.2439 | -35.3688 | -33.598
22.25 | -35.3692 | -33.6045
22.2561 | -35.3696 | -33.6109
22.2622 | -35.3701 | -33.6178
22.2683 | -35.3705 | -33.6242
22.2744 | -35.3709 | -33.631
22.2806 | -35.3713 | -33.6377
22.2867 | -35.3717 | -33.6442
22.2928 | -35.3722 | -33.6509
22.2989 | -35.3727 | -33.6582
22.305 | -35.373 | -33.665
22.3111 | -35.3735 | -33.6717
22.3172 | -35.3739 | -33.678
22.3233 | -35.3743 | -33.6846
22.3294 | -35.3748 | -33.6912
22.3355 | -35.3752 | -33.6976
22.3416 | -35.3756 | -33.7044
22.3477 | -35.376 | -33.7102
22.3538 | -35.3764 | -33.7174
22.3599 | -35.377 | -33.7239
22.366 | -35.3774 | -33.7304
22.3721 | -35.3778 | -33.7367
22.3782 | -35.3782 | -33.7435
22.3843 | -35.3786 | -33.7499
22.3904 | -35.3791 | -33.7566
22.3965 | -35.3795 | -33.763
22.4026 | -35.3799 | -33.7701
22.4087 | -35.3803 | -33.7761
22.4148 | -35.3807 | -33.7826
22.4209 | -35.3811 | -33.7895
22.4271 | -35.3815 | -33.7966
22.4332 | -35.3819 | -33.8032
22.4393 | -35.3823 | -33.8097
22.4454 | -35.3827 | -33.8161
22.4515 | -35.3831 | -33.8229
22.4576 | -35.3835 | -33.8294
22.4637 | -35.384 | -33.8363
22.4698 | -35.3844 | -33.8429
22.4759 | -35.3847 | -33.8499
22.482 | -35.385 | -33.8572
22.4881 | -35.3854 | -33.8638
22.4942 | -35.3859 | -33.8703
22.5003 | -35.3862 | -33.8768
22.5064 | -35.3867 | -33.8833
22.5125 | -35.3871 | -33.8893
22.5186 | -35.3876 | -33.8954
22.5247 | -35.3881 | -33.9017
22.5308 | -35.3885 | -33.9085
22.5369 | -35.3889 | -33.9155
22.543 | -35.3892 | -33.9223
22.5491 | -35.3897 | -33.9284
22.5552 | -35.3901 | -33.9354
22.5613 | -35.3906 | -33.9421
22.5675 | -35.391 | -33.949
22.5736 | -35.3914 | -33.9563
22.5797 | -35.3918 | -33.9634
22.5858 | -35.3923 | -33.9703
22.5919 | -35.3927 | -33.9767
22.598 | -35.3931 | -33.9836
22.6041 | -35.3935 | -33.9902
22.6102 | -35.394 | -33.9975
22.6163 | -35.3943 | -34.0038
22.6224 | -35.3947 | -34.0103
22.6285 | -35.3951 | -34.0167
22.6346 | -35.3955 | -34.0233
22.6407 | -35.3959 | -34.0298
22.6468 | -35.3964 | -34.0364
22.6529 | -35.3968 | -34.043
22.659 | -35.3972 | -34.0491
22.6651 | -35.3976 | -34.0557
22.6712 | -35.3981 | -34.0627
22.6773 | -35.3985 | -34.0698
22.6834 | -35.399 | -34.0759
22.6895 | -35.3994 | -34.0823
22.6956 | -35.3998 | -34.0888
22.7017 | -35.4002 | -34.0955
22.7079 | -35.4007 | -34.102
22.714 | -35.4011 | -34.1085
22.7201 | -35.4015 | -34.1152
22.7262 | -35.4018 | -34.1219
22.7323 | -35.4023 | -34.1288
22.7384 | -35.4027 | -34.1358
22.7445 | -35.4031 | -34.1425
22.7506 | -35.4035 | -34.1492
22.7567 | -35.4039 | -34.1559
22.7628 | -35.4043 | -34.1621
22.7689 | -35.4047 | -34.1689
22.775 | -35.4052 | -34.1756
22.7811 | -35.4056 | -34.1821
22.7872 | -35.4061 | -34.1884
22.7933 | -35.4065 | -34.1954
22.7994 | -35.4069 | -34.2022
22.8055 | -35.4074 | -34.2088
22.8116 | -35.4079 | -34.2154
22.8177 | -35.4083 | -34.2222
22.8238 | -35.4088 | -34.2287
22.8299 | -35.4092 | -34.2347
22.836 | -35.4096 | -34.2411
22.8421 | -35.41 | -34.2479
22.8482 | -35.4104 | -34.2538
22.8544 | -35.4108 | -34.2609
22.8605 | -35.4113 | -34.2677
22.8666 | -35.4115 | -34.2728
22.8727 | -35.412 | -34.2782
22.8788 | -35.4123 | -34.2848
22.8849 | -35.4128 | -34.2915
22.891 | -35.4132 | -34.2981
22.8971 | -35.4136 | -34.3046
22.9032 | -35.4139 | -34.3113
22.9093 | -35.4142 | -34.3179
22.9154 | -35.4147 | -34.3243
22.9215 | -35.4151 | -34.3307
22.9276 | -35.4154 | -34.3372
22.9337 | -35.4158 | -34.3439
22.9398 | -35.4162 | -34.3506
22.9459 | -35.4166 | -34.3568
22.952 | -35.4171 | -34.3634
22.9581 | -35.4175 | -34.3698
22.9642 | -35.4178 | -34.3761
22.9703 | -35.4182 | -34.3824
22.9764 | -35.4186 | -34.3888
22.9825 | -35.419 | -34.3951
22.9886 | -35.4194 | -34.4019
22.9948 | -35.4198 | -34.4084
23.0009 | -35.4202 | -34.4152
23.007 | -35.4206 | -34.4214
23.0131 | -35.4211 | -34.4281
23.0192 | -35.4215 | -34.4346
23.0253 | -35.4219 | -34.4412
23.0314 | -35.4223 | -34.4479
23.0375 | -35.4227 | -34.4544
23.0436 | -35.423 | -34.4609
23.0497 | -35.4234 | -34.4672
23.0558 | -35.4238 | -34.4734
23.0619 | -35.4241 | -34.4797
23.068 | -35.4245 | -34.4864
23.0741 | -35.4249 | -34.4921
23.0802 | -35.4253 | -34.4987
23.0863 | -35.4257 | -34.5053
23.0924 | -35.4261 | -34.5121
23.0985 | -35.4265 | -34.5188
23.1046 | -35.4269 | -34.5252
23.1107 | -35.4273 | -34.5324
23.1168 | -35.4276 | -34.5389
23.1229 | -35.4281 | -34.5454
23.129 | -35.4284 | -34.5515
23.1351 | -35.4288 | -34.5576
23.1413 | -35.4293 | -34.5644
23.1474 | -35.4296 | -34.571
23.1535 | -35.43 | -34.5777
23.1596 | -35.4304 | -34.5841
23.1657 | -35.4308 | -34.5905
23.1718 | -35.4312 | -34.5967
23.1779 | -35.4315 | -34.6032
23.184 | -35.432 | -34.6085
23.1901 | -35.4324 | -34.615
23.1962 | -35.4329 | -34.6213
23.2023 | -35.4333 | -34.6279
23.2084 | -35.4337 | -34.6344
23.2145 | -35.4342 | -34.6403
23.2206 | -35.4346 | -34.6464
23.2267 | -35.4351 | -34.6528
23.2328 | -35.4355 | -34.6596
23.2389 | -35.4359 | -34.666
23.245 | -35.4364 | -34.6727
23.2511 | -35.4367 | -34.6787
23.2572 | -35.4371 | -34.6849
23.2633 | -35.4375 | -34.6915
23.2694 | -35.4379 | -34.6978
23.2755 | -35.4382 | -34.7048
23.2817 | -35.4386 | -34.7117
23.2878 | -35.4391 | -34.7175
23.2939 | -35.4395 | -34.7237
23.3 | -35.4398 | -34.73
23.3061 | -35.4402 | -34.7365
23.3122 | -35.4406 | -34.743
23.3183 | -35.441 | -34.7492
23.3244 | -35.4414 | -34.7551
23.3305 | -35.4417 | -34.7615
23.3366 | -35.442 | -34.7676
23.3427 | -35.4424 | -34.7742
23.3488 | -35.4428 | -34.7804
23.3549 | -35.4432 | -34.7869
23.361 | -35.4437 | -34.7931
23.3671 | -35.444 | -34.7994
23.3732 | -35.4444 | -34.8059
23.3793 | -35.4449 | -34.8122
23.3854 | -35.4453 | -34.8185
23.3915 | -35.4457 | -34.8245
23.3976 | -35.4461 | -34.8307
23.4037 | -35.4465 | -34.837
23.4098 | -35.4469 | -34.8434
23.4159 | -35.4473 | -34.85
23.422 | -35.4477 | -34.8564
23.4282 | -35.4481 | -34.8625
23.4343 | -35.4485 | -34.8689
23.4404 | -35.4489 | -34.8754
23.4465 | -35.4492 | -34.8817
23.4526 | -35.4496 | -34.8878
23.4587 | -35.45 | -34.8942
23.4648 | -35.4503 | -34.9
23.4709 | -35.4507 | -34.9063
23.477 | -35.4511 | -34.9127
23.4831 | -35.4514 | -34.9193
23.4892 | -35.4518 | -34.9258
23.4953 | -35.4522 | -34.9321
23.5014 | -35.4525 | -34.9387
23.5075 | -35.4529 | -34.945
23.5136 | -35.4533 | -34.9515
23.5197 | -35.4537 | -34.9583
23.5258 | -35.4541 | -34.9647
23.5319 | -35.4545 | -34.9707
23.538 | -35.4549 | -34.9773
23.5441 | -35.4553 | -34.9839
23.5502 | -35.4557 | -34.99
23.5563 | -35.4561 | -34.9962
23.5624 | -35.4565 | -35.0028
23.5686 | -35.4569 | -35.0085
23.5747 | -35.4573 | -35.015
23.5808 | -35.4577 | -35.0212
23.5869 | -35.4581 | -35.0274
23.593 | -35.4585 | -35.0337
23.5991 | -35.4589 | -35.0401
23.6052 | -35.4593 | -35.0462
23.6113 | -35.4597 | -35.0526
23.6174 | -35.46 | -35.0587
23.6235 | -35.4604 | -35.065
23.6296 | -35.4608 | -35.0715
23.6357 | -35.4612 | -35.0777
23.6418 | -35.4616 | -35.0837
23.6479 | -35.462 | -35.0898
23.654 | -35.4623 | -35.0959
23.6601 | -35.4626 | -35.1024
23.6662 | -35.463 | -35.1085
23.6723 | -35.4633 | -35.1142
23.6784 | -35.4637 | -35.1203
23.6845 | -35.464 | -35.1264
23.6906 | -35.4644 | -35.1329
23.6967 | -35.4648 | -35.1392
23.7028 | -35.4652 | -35.1457
23.7089 | -35.4655 | -35.152
23.7151 | -35.4659 | -35.1583
23.7212 | -35.4662 | -35.1646
23.7273 | -35.4666 | -35.1709
23.7334 | -35.467 | -35.1771
23.7395 | -35.4674 | -35.1834
23.7456 | -35.4677 | -35.189
23.7517 | -35.4681 | -35.1954
23.7578 | -35.4684 | -35.2016
23.7639 | -35.4688 | -35.2078
23.77 | -35.4692 | -35.214
23.7761 | -35.4696 | -35.2207
23.7822 | -35.47 | -35.2266
23.7883 | -35.4704 | -35.2328
23.7944 | -35.4708 | -35.239
23.8005 | -35.4711 | -35.2453
23.8066 | -35.4716 | -35.2517
23.8127 | -35.4719 | -35.2574
23.8188 | -35.4723 | -35.2638
23.8249 | -35.4727 | -35.2701
23.831 | -35.473 | -35.2761
23.8371 | -35.4734 | -35.2823
23.8432 | -35.4738 | -35.2886
23.8493 | -35.4742 | -35.2946
23.8555 | -35.4746 | -35.3009
23.8616 | -35.475 | -35.307
23.8677 | -35.4754 | -35.3133
23.8738 | -35.4757 | -35.3195
23.8799 | -35.4761 | -35.3257
23.886 | -35.4765 | -35.3316
23.8921 | -35.4769 | -35.3377
23.8982 | -35.4773 | -35.3438
23.9043 | -35.4777 | -35.3499
23.9104 | -35.478 | -35.3561
23.9165 | -35.4784 | -35.3622
23.9226 | -35.4788 | -35.3681
23.9287 | -35.4791 | -35.3744
23.9348 | -35.4795 | -35.3805
23.9409 | -35.4799 | -35.3866
23.947 | -35.4803 | -35.3925
23.9531 | -35.4806 | -35.3984
23.9592 | -35.481 | -35.4045
23.9653 | -35.4814 | -35.4107
23.9714 | -35.4817 | -35.417
23.9775 | -35.4821 | -35.4233
23.9836 | -35.4825 | -35.4296
23.9897 | -35.4829 | -35.4354
23.9958 | -35.4833 | -35.4415
24.002 | -35.4836 | -35.4477
24.0081 | -35.484 | -35.4539
24.0142 | -35.4844 | -35.4598
24.0203 | -35.4848 | -35.466
24.0264 | -35.4852 | -35.4717
24.0325 | -35.4856 | -35.4778
24.0386 | -35.486 | -35.484
24.0447 | -35.4864 | -35.4903
24.0508 | -35.4867 | -35.4965
24.0569 | -35.4871 | -35.5027
24.063 | -35.4875 | -35.5082
24.0691 | -35.4879 | -35.5144
24.0752 | -35.4883 | -35.5206
24.0813 | -35.4886 | -35.5267
24.0874 | -35.4891 | -35.5328
24.0935 | -35.4894 | -35.539
24.0996 | -35.4898 | -35.5449
24.1057 | -35.4902 | -35.5512
24.1118 | -35.4906 | -35.5575
24.1179 | -35.491 | -35.5637
24.124 | -35.4913 | -35.5697
24.1301 | -35.4917 | -35.5754
24.1362 | -35.4921 | -35.5814
24.1424 | -35.4925 | -35.5878
24.1485 | -35.4929 | -35.5942
24.1546 | -35.4933 | -35.6002
24.1607 | -35.4936 | -35.6061
24.1668 | -35.494 | -35.6122
24.1729 | -35.4943 | -35.6184
24.179 | -35.4947 | -35.6246
24.1851 | -35.4951 | -35.6307
24.1912 | -35.4955 | -35.6367
24.1973 | -35.4958 | -35.6428
24.2034 | -35.4962 | -35.6488
24.2095 | -35.4966 | -35.6553
24.2156 | -35.497 | -35.6614
24.2217 | -35.4974 | -35.6675
24.2278 | -35.4978 | -35.6735
24.2339 | -35.4981 | -35.6795
24.24 | -35.4985 | -35.6852
24.2461 | -35.4989 | -35.6916
24.2522 | -35.4993 | -35.6979
24.2583 | -35.4997 | -35.7039
24.2644 | -35.5001 | -35.71
24.2705 | -35.5005 | -35.7158
24.2766 | -35.5009 | -35.7218
24.2827 | -35.5012 | -35.7282
24.2889 | -35.5016 | -35.7345
24.295 | -35.502 | -35.7406
24.3011 | -35.5024 | -35.7463
24.3072 | -35.5028 | -35.752
24.3133 | -35.5031 | -35.7582
24.3194 | -35.5035 | -35.7644
24.3255 | -35.5039 | -35.7706
24.3316 | -35.5042 | -35.7768
24.3377 | -35.5046 | -35.7828
24.3438 | -35.505 | -35.7882
24.3499 | -35.5053 | -35.7945
24.356 | -35.5057 | -35.8004
24.3621 | -35.506 | -35.8066
24.3682 | -35.5064 | -35.8126
24.3743 | -35.5068 | -35.8186
24.3804 | -35.5072 | -35.8242
24.3865 | -35.5076 | -35.8303
24.3926 | -35.508 | -35.8363
24.3987 | -35.5084 | -35.8425
24.4048 | -35.5087 | -35.8485
24.4109 | -35.5091 | -35.8542
24.417 | -35.5095 | -35.86
24.4231 | -35.5099 | -35.8659
24.4293 | -35.5103 | -35.8717
24.4354 | -35.5106 | -35.8776
24.4415 | -35.511 | -35.8837
24.4476 | -35.5114 | -35.8898
24.4537 | -35.5119 | -35.8959
24.4598 | -35.5123 | -35.9018
24.4659 | -35.5126 | -35.908
24.472 | -35.513 | -35.9138
24.4781 | -35.5134 | -35.9195
24.4842 | -35.5138 | -35.925
24.4903 | -35.5142 | -35.9309
24.4964 | -35.5145 | -35.9369
24.5025 | -35.5149 | -35.9428
24.5086 | -35.5153 | -35.9488
24.5147 | -35.5157 | -35.9547
24.5208 | -35.5161 | -35.9603
24.5269 | -35.5165 | -35.9664
24.533 | -35.5168 | -35.9724
24.5391 | -35.5173 | -35.9784
24.5452 | -35.5176 | -35.9845
24.5513 | -35.518 | -35.9901
24.5574 | -35.5184 | -35.9959
24.5635 | -35.5188 | -36.0017
24.5696 | -35.5192 | -36.0077
24.5758 | -35.5196 | -36.0136
24.5819 | -35.52 | -36.0197
24.588 | -35.5204 | -36.0253
24.5941 | -35.5207 | -36.0312
24.6002 | -35.5212 | -36.0371
24.6063 | -35.5215 | -36.0432
24.6124 | -35.5219 | -36.0493
24.6185 | -35.5223 | -36.0553
24.6246 | -35.5228 | -36.061
24.6307 | -35.5231 | -36.0669
24.6368 | -35.5235 | -36.0727
24.6429 | -35.5239 | -36.0789
24.649 | -35.5243 | -36.0851
24.6551 | -35.5247 | -36.0909
24.6612 | -35.5251 | -36.0965
24.6673 | -35.5256 | -36.1025
24.6734 | -35.526 | -36.1086
24.6795 | -35.5264 | -36.1148
24.6856 | -35.5268 | -36.1205
24.6917 | -35.5272 | -36.1262
24.6978 | -35.5275 | -36.1318
24.7039 | -35.5279 | -36.1377
24.71 | -35.5283 | -36.1436
24.7162 | -35.5287 | -36.1495
24.7223 | -35.5291 | -36.1554
24.7284 | -35.5296 | -36.1609
24.7345 | -35.53 | -36.1665
24.7406 | -35.5304 | -36.1724
24.7467 | -35.5308 | -36.1784
24.7528 | -35.5312 | -36.1841
24.7589 | -35.5315 | -36.1899
24.765 | -35.532 | -36.1952
24.7711 | -35.5324 | -36.2008
24.7772 | -35.5327 | -36.2065
24.7833 | -35.5331 | -36.2124
24.7894 | -35.5335 | -36.2184
24.7955 | -35.5339 | -36.2242
24.8016 | -35.5343 | -36.2296
24.8077 | -35.5347 | -36.2354
24.8138 | -35.5351 | -36.2414
24.8199 | -35.5355 | -36.2474
24.826 | -35.5359 | -36.2532
24.8321 | -35.5362 | -36.2591
24.8382 | -35.5367 | -36.2645
24.8443 | -35.5371 | -36.2703
24.8504 | -35.5375 | -36.2761
24.8565 | -35.5379 | -36.2821
24.8627 | -35.5382 | -36.2881
24.8688 | -35.5387 | -36.2937
24.8749 | -35.539 | -36.2996
24.881 | -35.5394 | -36.3053
24.8871 | -35.5398 | -36.3115
24.8932 | -35.5402 | -36.3174
24.8993 | -35.5406 | -36.3232
24.9054 | -35.541 | -36.3286
24.9115 | -35.5415 | -36.3344
24.9176 | -35.5419 | -36.3404
24.9237 | -35.5423 | -36.3459
24.9298 | -35.5427 | -36.3517
24.9359 | -35.5431 | -36.3574
24.942 | -35.5435 | -36.3626
24.9481 | -35.5439 | -36.3688
24.9542 | -35.5442 | -36.3746
24.9603 | -35.5446 | -36.3803
24.9664 | -35.545 | -36.3861
24.9725 | -35.5454 | -36.3921
24.9786 | -35.5458 | -36.3975
24.9847 | -35.5462 | -36.4033
24.9908 | -35.5466 | -36.409
24.9969 | -35.5469 | -36.415
25.0031 | -35.5473 | -36.4207
25.0092 | -35.5478 | -36.4262
25.0153 | -35.5482 | -36.4319
25.0214 | -35.5486 | -36.4378
25.0275 | -35.5489 | -36.4436
25.0336 | -35.5493 | -36.4495
25.0397 | -35.5497 | -36.4552
25.0458 | -35.5501 | -36.4606
25.0519 | -35.5505 | -36.4665
25.058 | -35.5509 | -36.4722
25.0641 | -35.5513 | -36.4779
25.0702 | -35.5517 | -36.4835
25.0763 | -35.5521 | -36.489
25.0824 | -35.5525 | -36.4943
25.0885 | -35.5528 | -36.5003
25.0946 | -35.5532 | -36.5059
25.1007 | -35.5536 | -36.5117
25.1068 | -35.554 | -36.5174
25.1129 | -35.5544 | -36.5231
25.119 | -35.5549 | -36.5286
25.1251 | -35.5553 | -36.5345
25.1312 | -35.5557 | -36.5401
25.1373 | -35.5561 | -36.546
25.1435 | -35.5565 | -36.5516
25.1496 | -35.5569 | -36.557
25.1557 | -35.5573 | -36.5627
25.1618 | -35.5577 | -36.5686
25.1679 | -35.5581 | -36.5742
25.174 | -35.5585 | -36.58
25.1801 | -35.5589 | -36.5857
25.1862 | -35.5593 | -36.5907
25.1923 | -35.5597 | -36.5964
25.1984 | -35.5601 | -36.602
25.2045 | -35.5605 | -36.6077
25.2106 | -35.5609 | -36.6134
25.2167 | -35.5613 | -36.6192
25.2228 | -35.5617 | -36.6251
25.2289 | -35.562 | -36.6308
25.235 | -35.5625 | -36.6369
25.2411 | -35.5629 | -36.6425
25.2472 | -35.5633 | -36.6484
25.2533 | -35.5637 | -36.6541
25.2594 | -35.5641 | -36.6597
25.2655 | -35.5645 | -36.6654
25.2716 | -35.5649 | -36.671
25.2777 | -35.5653 | -36.6767
25.2838 | -35.5657 | -36.6825
25.29 | -35.5661 | -36.6884
25.2961 | -35.5665 | -36.6937
25.3022 | -35.567 | -36.6991
25.3083 | -35.5673 | -36.7048
25.3144 | -35.5677 | -36.7103
25.3205 | -35.5681 | -36.7159
25.3266 | -35.5685 | -36.7212
25.3327 | -35.5689 | -36.7271
25.3388 | -35.5692 | -36.7325
25.3449 | -35.5696 | -36.7377
25.351 | -35.57 | -36.7436
25.3571 | -35.5704 | -36.7488
25.3632 | -35.5708 | -36.7541
25.3693 | -35.5712 | -36.7596
25.3754 | -35.5716 | -36.7652
25.3815 | -35.572 | -36.7714
25.3876 | -35.5723 | -36.7768
25.3937 | -35.5727 | -36.7824
25.3998 | -35.5731 | -36.7874
25.4059 | -35.5736 | -36.7932
25.412 | -35.574 | -36.7989
25.4181 | -35.5744 | -36.8044
25.4242 | -35.5748 | -36.8102
25.4304 | -35.5752 | -36.8159
25.4365 | -35.5757 | -36.8216
25.4426 | -35.5761 | -36.8272
25.4487 | -35.5764 | -36.8325
25.4548 | -35.5768 | -36.8384
25.4609 | -35.5772 | -36.8439
25.467 | -35.5776 | -36.8497
25.4731 | -35.578 | -36.8549
25.4792 | -35.5784 | -36.8608
25.4853 | -35.5788 | -36.8665
25.4914 | -35.5792 | -36.8721
25.4975 | -35.5795 | -36.8779
25.5036 | -35.5799 | -36.8834
25.5097 | -35.5803 | -36.8887
25.5158 | -35.5807 | -36.8947
25.5219 | -35.5811 | -36.9006
25.528 | -35.5815 | -36.9061
25.5341 | -35.582 | -36.9116
25.5402 | -35.5824 | -36.9171
25.5463 | -35.5828 | -36.9228
25.5524 | -35.5832 | -36.9285
25.5585 | -35.5836 | -36.9341
25.5646 | -35.5841 | -36.9397
25.5707 | -35.5845 | -36.9451
25.5769 | -35.5849 | -36.9504
25.583 | -35.5853 | -36.9561
25.5891 | -35.5857 | -36.9614
25.5952 | -35.5861 | -36.9671
25.6013 | -35.5865 | -36.9723
25.6074 | -35.5869 | -36.977
25.6135 | -35.5873 | -36.9827
25.6196 | -35.5877 | -36.9883
25.6257 | -35.5881 | -36.9941
25.6318 | -35.5885 | -36.9995
25.6379 | -35.5889 | -37.0051
25.644 | -35.5893 | -37.0104
25.6501 | -35.5896 | -37.0158
25.6562 | -35.5901 | -37.0215
25.6623 | -35.5904 | -37.0271
25.6684 | -35.5909 | -37.0327
25.6745 | -35.5913 | -37.0382
25.6806 | -35.5917 | -37.0432
25.6867 | -35.5922 | -37.0488
25.6928 | -35.5926 | -37.054
25.6989 | -35.593 | -37.0597
25.705 | -35.5934 | -37.0652
25.7111 | -35.5938 | -37.0709
25.7173 | -35.5942 | -37.0757
25.7234 | -35.5946 | -37.0815
25.7295 | -35.595 | -37.087
25.7356 | -35.5953 | -37.0928
25.7417 | -35.5957 | -37.0992
25.7478 | -35.5961 | -37.1043
25.7539 | -35.5965 | -37.1099
25.76 | -35.5969 | -37.1154
25.7661 | -35.5973 | -37.1211
25.7722 | -35.5977 | -37.1267
25.7783 | -35.5981 | -37.1323
25.7844 | -35.5984 | -37.1373
25.7905 | -35.5989 | -37.1428
25.7966 | -35.5993 | -37.1488
25.8027 | -35.5997 | -37.1548
25.8088 | -35.6001 | -37.1604
25.8149 | -35.6005 | -37.1659
25.821 | -35.601 | -37.1712
25.8271 | -35.6014 | -37.1765
25.8332 | -35.6018 | -37.1819
25.8393 | -35.6022 | -37.1879
25.8454 | -35.6026 | -37.1936
25.8515 | -35.603 | -37.1991
25.8576 | -35.6035 | -37.2046
25.8638 | -35.6039 | -37.21
25.8699 | -35.6043 | -37.2158
25.876 | -35.6047 | -37.2211
25.8821 | -35.6051 | -37.2267
25.8882 | -35.6055 | -37.2321
25.8943 | -35.6059 | -37.2375
25.9004 | -35.6063 | -37.2431
25.9065 | -35.6067 | -37.2486
25.9126 | -35.607 | -37.2541
25.9187 | -35.6075 | -37.2593
25.9248 | -35.6079 | -37.2643
25.9309 | -35.6083 | -37.27
25.937 | -35.6087 | -37.276
25.9431 | -35.6091 | -37.2817
25.9492 | -35.6095 | -37.2873
25.9553 | -35.6099 | -37.293
25.9614 | -35.6103 | -37.2984
25.9675 | -35.6107 | -37.3038
25.9736 | -35.6111 | -37.309
25.9797 | -35.6115 | -37.3141
25.9858 | -35.6118 | -37.3195
25.9919 | -35.6123 | -37.3244
25.998 | -35.6127 | -37.3299
26.0042 | -35.6131 | -37.3352
26.0103 | -35.6134 | -37.3405
26.0164 | -35.6138 | -37.346
26.0225 | -35.6142 | -37.3518
26.0286 | -35.6146 | -37.3572
26.0347 | -35.615 | -37.3622
26.0408 | -35.6154 | -37.3676
26.0469 | -35.6159 | -37.3729
26.053 | -35.6163 | -37.3788
26.0591 | -35.6167 | -37.3847
26.0652 | -35.6171 | -37.3899
26.0713 | -35.6175 | -37.3954
26.0774 | -35.6179 | -37.4007
26.0835 | -35.6183 | -37.4063
26.0896 | -35.6188 | -37.4115
26.0957 | -35.6191 | -37.417
26.1018 | -35.6194 | -37.4224
26.1079 | -35.6199 | -37.4278
26.114 | -35.6203 | -37.4332
26.1201 | -35.6207 | -37.4389
26.1262 | -35.6211 | -37.4444
26.1323 | -35.6215 | -37.4501
26.1384 | -35.6219 | -37.4546
26.1445 | -35.6223 | -37.46
26.1507 | -35.6227 | -37.4657
26.1568 | -35.6231 | -37.471
26.1629 | -35.6235 | -37.4765
26.169 | -35.6239 | -37.4817
26.1751 | -35.6242 | -37.4865
26.1812 | -35.6246 | -37.4917
26.1873 | -35.625 | -37.4973
26.1934 | -35.6254 | -37.5026
26.1995 | -35.6258 | -37.5081
26.2056 | -35.6263 | -37.5134
26.2117 | -35.6267 | -37.5188
26.2178 | -35.627 | -37.5242
26.2239 | -35.6274 | -37.5293
26.23 | -35.6279 | -37.535
26.2361 | -35.6283 | -37.5405
26.2422 | -35.6287 | -37.5448
26.2483 | -35.6291 | -37.5504
26.2544 | -35.6295 | -37.5553
26.2605 | -35.6299 | -37.5608
26.2666 | -35.6303 | -37.5664
26.2727 | -35.6307 | -37.5717
26.2788 | -35.6311 | -37.5769
26.2849 | -35.6315 | -37.5823
26.2911 | -35.6319 | -37.5882
26.2972 | -35.6324 | -37.5932
26.3033 | -35.6327 | -37.5985
26.3094 | -35.6332 | -37.6037
26.3155 | -35.6336 | -37.6088
26.3216 | -35.634 | -37.6141
26.3277 | -35.6344 | -37.6191
26.3338 | -35.6347 | -37.6248
26.3399 | -35.6351 | -37.6297
26.346 | -35.6355 | -37.6354
26.3521 | -35.6359 | -37.6406
26.3582 | -35.6364 | -37.6461
26.3643 | -35.6368 | -37.6514
26.3704 | -35.6372 | -37.6567
26.3765 | -35.6376 | -37.6622
26.3826 | -35.6379 | -37.6671
26.3887 | -35.6383 | -37.6726
26.3948 | -35.6386 | -37.6778
26.4009 | -35.6391 | -37.6831
26.407 | -35.6395 | -37.6885
26.4131 | -35.6399 | -37.694
26.4192 | -35.6402 | -37.6987
26.4253 | -35.6406 | -37.7039
26.4314 | -35.641 | -37.7091
26.4376 | -35.6413 | -37.7147
26.4437 | -35.6417 | -37.7205
26.4498 | -35.6421 | -37.7262
26.4559 | -35.6425 | -37.7317
26.462 | -35.6429 | -37.737
26.4681 | -35.6432 | -37.7426
26.4742 | -35.6437 | -37.7481
26.4803 | -35.6441 | -37.7536
26.4864 | -35.6446 | -37.758
26.4925 | -35.645 | -37.7631
26.4986 | -35.6454 | -37.7681
26.5047 | -35.6458 | -37.7734
26.5108 | -35.6463 | -37.7788
26.5169 | -35.6467 | -37.7841
26.523 | -35.647 | -37.7894
26.5291 | -35.6474 | -37.7947
26.5352 | -35.6478 | -37.7998
26.5413 | -35.6482 | -37.8056
26.5474 | -35.6487 | -37.811
26.5535 | -35.649 | -37.8162
26.5596 | -35.6494 | -37.8215
26.5657 | -35.6497 | -37.8269
26.5718 | -35.6502 | -37.8324
26.578 | -35.6506 | -37.8376
26.5841 | -35.651 | -37.843
26.5902 | -35.6514 | -37.8485
26.5963 | -35.6518 | -37.8535
26.6024 | -35.6521 | -37.8591
26.6085 | -35.6525 | -37.8645
26.6146 | -35.6529 | -37.8703
26.6207 | -35.6532 | -37.8756
26.6268 | -35.6536 | -37.8813
26.6329 | -35.6541 | -37.8863
26.639 | -35.6544 | -37.8918
26.6451 | -35.6548 | -37.8971
26.6512 | -35.6552 | -37.9024
26.6573 | -35.6556 | -37.9075
26.6634 | -35.656 | -37.9131
26.6695 | -35.6563 | -37.918
26.6756 | -35.6568 | -37.9234
26.6817 | -35.6572 | -37.929
26.6878 | -35.6576 | -37.9342
26.6939 | -35.658 | -37.9396
26.7 | -35.6584 | -37.945
26.7061 | -35.6587 | -37.9505
26.7122 | -35.6591 | -37.9558
26.7183 | -35.6594 | -37.961
26.7245 | -35.6597 | -37.9661
26.7306 | -35.6601 | -37.971
26.7367 | -35.6605 | -37.9757
26.7428 | -35.6608 | -37.9809
26.7489 | -35.6612 | -37.9861
26.755 | -35.6615 | -37.9911
26.7611 | -35.6619 | -37.9963
26.7672 | -35.6622 | -38.0016
26.7733 | -35.6626 | -38.0073
26.7794 | -35.663 | -38.0129
26.7855 | -35.6635 | -38.0177
26.7916 | -35.6638 | -38.0228
26.7977 | -35.6643 | -38.0279
26.8038 | -35.6646 | -38.0327
26.8099 | -35.6651 | -38.0387
26.816 | -35.6655 | -38.0439
26.8221 | -35.6659 | -38.0492
26.8282 | -35.6663 | -38.0545
26.8343 | -35.6667 | -38.0592
26.8404 | -35.6672 | -38.0634
26.8465 | -35.6676 | -38.0686
26.8526 | -35.668 | -38.0742
26.8587 | -35.6685 | -38.0797
26.8649 | -35.6689 | -38.0843
26.871 | -35.6694 | -38.0903
26.8771 | -35.6697 | -38.0948
26.8832 | -35.6702 | -38.0998
26.8893 | -35.6708 | -38.1056
26.8954 | -35.6711 | -38.1113
26.9015 | -35.6716 | -38.117
26.9076 | -35.672 | -38.1222
26.9137 | -35.6723 | -38.1271
26.9198 | -35.6726 | -38.1329
26.9259 | -35.6731 | -38.1385
26.932 | -35.6734 | -38.1435
26.9381 | -35.6739 | -38.1485
26.9442 | -35.6743 | -38.1536
26.9503 | -35.6747 | -38.1589
26.9564 | -35.6751 | -38.1642
26.9625 | -35.6755 | -38.1691
26.9686 | -35.676 | -38.1739
26.9747 | -35.6763 | -38.1791
26.9808 | -35.6767 | -38.1842
26.9869 | -35.677 | -38.19
26.993 | -35.6775 | -38.1953
26.9991 | -35.6778 | -38.2007
27.0052 | -35.6782 | -38.2057
27.0114 | -35.6786 | -38.2111
27.0175 | -35.679 | -38.217
27.0236 | -35.6794 | -38.2225
27.0297 | -35.6798 | -38.2277
27.0358 | -35.6802 | -38.2331
27.0419 | -35.6804 | -38.2382
27.048 | -35.6808 | -38.2436
27.0541 | -35.6812 | -38.2485
27.0602 | -35.6816 | -38.2541
27.0663 | -35.6821 | -38.2595
27.0724 | -35.6825 | -38.265
27.0785 | -35.6829 | -38.2703
27.0846 | -35.6833 | -38.2752
27.0907 | -35.6837 | -38.2804
27.0968 | -35.6841 | -38.2858
27.1029 | -35.6845 | -38.291
27.109 | -35.6849 | -38.2963
27.1151 | -35.6853 | -38.3006
27.1212 | -35.6856 | -38.3052
27.1273 | -35.686 | -38.3105
27.1334 | -35.6864 | -38.3161
27.1395 | -35.6867 | -38.3211
27.1456 | -35.6871 | -38.326
27.1518 | -35.6875 | -38.3314
27.1579 | -35.6879 | -38.3377
27.164 | -35.6882 | -38.3433
27.1701 | -35.6887 | -38.3483
27.1762 | -35.6891 | -38.3535
27.1823 | -35.6895 | -38.3588
27.1884 | -35.6898 | -38.3642
27.1945 | -35.6903 | -38.3689
27.2006 | -35.6907 | -38.3737
27.2067 | -35.6913 | -38.3789
27.2128 | -35.6917 | -38.3841
27.2189 | -35.6921 | -38.389
27.225 | -35.6925 | -38.3946
27.2311 | -35.6929 | -38.4
27.2372 | -35.6933 | -38.4056
27.2433 | -35.6937 | -38.4107
27.2494 | -35.694 | -38.4155
27.2555 | -35.6944 | -38.4213
27.2616 | -35.6948 | -38.4266
27.2677 | -35.6951 | -38.4318
27.2738 | -35.6954 | -38.4369
27.2799 | -35.6958 | -38.4423
27.286 | -35.6961 | -38.4479
27.2921 | -35.6965 | -38.4529
27.2983 | -35.6968 | -38.458
27.3044 | -35.6972 | -38.4631
27.3105 | -35.6976 | -38.4683
27.3166 | -35.698 | -38.4738
27.3227 | -35.6983 | -38.4787
27.3288 | -35.6987 | -38.4842
27.3349 | -35.699 | -38.4893
27.341 | -35.6994 | -38.4947
27.3471 | -35.6998 | -38.5
27.3532 | -35.7001 | -38.505
27.3593 | -35.7005 | -38.5103
27.3654 | -35.701 | -38.5155
27.3715 | -35.7013 | -38.5203
27.3776 | -35.7017 | -38.5257
27.3837 | -35.7022 | -38.5311
27.3898 | -35.7026 | -38.536
27.3959 | -35.7029 | -38.5414
27.402 | -35.7033 | -38.5465
27.4081 | -35.7037 | -38.5517
27.4142 | -35.7041 | -38.5569
27.4203 | -35.7044 | -38.562
27.4264 | -35.7047 | -38.567
27.4325 | -35.7051 | -38.5725
27.4387 | -35.7054 | -38.5778
27.4448 | -35.7057 | -38.5825
27.4509 | -35.7062 | -38.5877
27.457 | -35.7066 | -38.593
27.4631 | -35.707 | -38.5981
27.4692 | -35.7074 | -38.6033
27.4753 | -35.7078 | -38.6084
27.4814 | -35.7082 | -38.6134
27.4875 | -35.7086 | -38.6184
27.4936 | -35.709 | -38.624
27.4997 | -35.7093 | -38.6292
27.5058 | -35.7097 | -38.6342
27.5119 | -35.7101 | -38.6395
27.518 | -35.7105 | -38.6448
27.5241 | -35.7109 | -38.65
27.5302 | -35.7112 | -38.6552
27.5363 | -35.7116 | -38.6607
27.5424 | -35.7119 | -38.6665
27.5485 | -35.7123 | -38.6718
27.5546 | -35.7127 | -38.6771
27.5607 | -35.713 | -38.6823
27.5668 | -35.7134 | -38.6877
27.5729 | -35.7137 | -38.6927
27.5791 | -35.714 | -38.6977
27.5852 | -35.7144 | -38.7029
27.5913 | -35.7147 | -38.7084
27.5974 | -35.715 | -38.714
27.6035 | -35.7153 | -38.7195
27.6096 | -35.7157 | -38.7249
27.6157 | -35.7161 | -38.7301
27.6218 | -35.7164 | -38.7353
27.6279 | -35.7167 | -38.7407
27.634 | -35.7171 | -38.7458
27.6401 | -35.7175 | -38.7506
27.6462 | -35.7178 | -38.756
27.6523 | -35.7181 | -38.7607
27.6584 | -35.7185 | -38.766
27.6645 | -35.7189 | -38.7714
27.6706 | -35.7193 | -38.7768
27.6767 | -35.7196 | -38.7817
27.6828 | -35.72 | -38.7872
27.6889 | -35.7203 | -38.7922
27.695 | -35.7207 | -38.7976
27.7011 | -35.7211 | -38.8028
27.7072 | -35.7214 | -38.8077
27.7133 | -35.7218 | -38.813
27.7194 | -35.7222 | -38.818
27.7256 | -35.7226 | -38.8233
27.7317 | -35.7229 | -38.8283
27.7378 | -35.7233 | -38.8339
27.7439 | -35.7237 | -38.8396
27.75 | -35.7241 | -38.8451
27.7561 | -35.7244 | -38.8502
27.7622 | -35.7248 | -38.8552
27.7683 | -35.7252 | -38.8602
27.7744 | -35.7256 | -38.8656
27.7805 | -35.7259 | -38.8707
27.7866 | -35.7263 | -38.876
27.7927 | -35.7267 | -38.8809
27.7988 | -35.7271 | -38.8861
27.8049 | -35.7274 | -38.8913
27.811 | -35.7278 | -38.8965
27.8171 | -35.7282 | -38.9022
27.8232 | -35.7285 | -38.9069
27.8293 | -35.7289 | -38.9126
27.8354 | -35.7292 | -38.9178
27.8415 | -35.7296 | -38.9227
27.8476 | -35.7299 | -38.9276
27.8537 | -35.7304 | -38.9327
27.8598 | -35.7308 | -38.9375
27.866 | -35.7311 | -38.9432
27.8721 | -35.7315 | -38.9486
27.8782 | -35.7318 | -38.9536
27.8843 | -35.7322 | -38.9588
27.8904 | -35.7326 | -38.9638
27.8965 | -35.733 | -38.9696
27.9026 | -35.7334 | -38.9751
27.9087 | -35.7337 | -38.9801
27.9148 | -35.7341 | -38.985
27.9209 | -35.7345 | -38.9902
27.927 | -35.7348 | -38.9952
27.9331 | -35.7351 | -39.0001
27.9392 | -35.7355 | -39.0047
27.9453 | -35.7359 | -39.0098
27.9514 | -35.7362 | -39.0151
27.9575 | -35.7365 | -39.0205
27.9636 | -35.7368 | -39.0259
27.9697 | -35.7372 | -39.0309
27.9758 | -35.7376 | -39.0363
27.9819 | -35.738 | -39.0413
27.988 | -35.7384 | -39.0464
27.9941 | -35.7388 | -39.0518
28.0002 | -35.7391 | -39.057
28.0063 | -35.7394 | -39.0622
28.0125 | -35.7398 | -39.0666
28.0186 | -35.7401 | -39.0717
28.0247 | -35.7405 | -39.0768
28.0308 | -35.7408 | -39.0821
28.0369 | -35.7411 | -39.0866
28.043 | -35.7414 | -39.0918
28.0491 | -35.7418 | -39.0971
28.0552 | -35.7422 | -39.1023
28.0613 | -35.7426 | -39.1076
28.0674 | -35.7429 | -39.1128
28.0735 | -35.7432 | -39.1177

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.5005 | -44.5137
21.4908 | -35.5491 | -36.1887
16.4948 | -34.9323 | -28.5711
12.6602 | -34.5683 | -22.1253
9.7171 | -34.3398 | -17.0908
7.4582 | -34.1818 | -13.0104
5.7244 | -34.0763 | -9.7015
4.3936 | -34.015 | -7.0424
3.3722 | -33.9911 | -4.9476
2.5883 | -33.9785 | -3.3493
1.9866 | -33.975 | -2.0829
1.5248 | -33.9592 | -1.06
1.1703 | -33.9711 | -0.39484
1 | -33.956 | -0.065612
0.89822 | -33.9457 | 0.2126
0.68941 | -33.9348 | 0.8073
0.52919 | -33.9485 | 1.269
0.40612 | -33.939 | 1.8824
0.31171 | -33.9294 | 2.3415
0.23928 | -33.9403 | 3.0749
0.18358 | -33.9219 | 3.5087
0.14091 | -33.9374 | 4.3932
0.10819 | -33.9572 | 5.5705
0.083017 | -34.0112 | 7.0438
0.063705 | -34.0507 | 8.9307
0.048923 | -34.0998 | 11.4386
0.037479 | -34.1984 | 14.6181
0.028801 | -34.385 | 18.6298
0.022078 | -34.7022 | 23.9578
0.017023 | -35.1729 | 30.3315
0.013018 | -35.992 | 38.0755
0.010014 | -37.1013 | 45.9695

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.00366 | -11.9412
0.00732 | -29.8067
0.01098 | -35.9884
0.014641 | -40.7683
0.018301 | -42.8973
0.021961 | -44.0488
0.025621 | -45.8307
0.029281 | -47.0643
0.032941 | -47.0221
0.036602 | -48.5382
0.040262 | -50.5114
0.043922 | -52.0085
0.047582 | -53.8052
0.051242 | -55.4213
0.054902 | -54.8807
0.058562 | -55.639
0.062223 | -57.8832
0.065883 | -57.3096
0.069543 | -57.3648
0.073203 | -58.0445
0.076863 | -59.2935
0.080523 | -58.4166
0.084184 | -57.6998
0.087844 | -59.6154
0.091504 | -61.7056
0.095164 | -62.7549
0.098824 | -64.1934
0.10248 | -64.3232
0.10614 | -63.7038
0.1098 | -62.505
0.11346 | -63.2281
0.11713 | -64.2899
0.12079 | -64.8374
0.12445 | -64.8243
0.1281 | -66.0092
0.13177 | -66.3707
0.13543 | -65.7692
0.13909 | -66.1946
0.14275 | -65.8536
0.14641 | -66.2878
0.15007 | -67.0219
0.15373 | -67.004
0.15739 | -67.6784
0.16105 | -67.8943
0.16471 | -67.2641
0.16837 | -68.4744
0.17203 | -68.4649
0.17569 | -68.1281
0.17935 | -68.5085
0.18301 | -68.1314
0.18667 | -67.9414
0.19033 | -69.4536
0.19399 | -70.3053
0.19765 | -68.9292
0.20131 | -68.1441
0.20497 | -69.5319
0.20863 | -71.5143
0.21229 | -70.3491
0.21595 | -70.2667
0.21961 | -70.6061
0.22327 | -70.5518
0.22693 | -71.8486
0.23059 | -72.4222
0.23425 | -72.6652
0.23791 | -73.3687
0.24157 | -73.3515
0.24523 | -71.9825
0.24889 | -72.4308
0.25255 | -73.0274
0.25621 | -72.5266
0.25987 | -72.8914
0.26353 | -72.8194
0.26719 | -72.8903
0.27085 | -73.258
0.27451 | -74.2118
0.27817 | -74.205
0.28183 | -74.7429
0.28549 | -74.0577
0.28915 | -73.8747
0.29281 | -73.8706
0.29647 | -73.921
0.30013 | -73.7117
0.30379 | -72.9701
0.30745 | -73.7535
0.31111 | -74.784
0.31477 | -74.1295
0.31843 | -74.1963
0.32209 | -74.7171
0.32575 | -75.9824
0.32941 | -74.9352
0.33307 | -74.2546
0.33673 | -74.97
0.34039 | -75.0355
0.34405 | -74.7443
0.34771 | -74.7759
0.35137 | -75.8972
0.35503 | -76.3213
0.35869 | -75.6437
0.36235 | -76.4384
0.36601 | -76.4797
0.36968 | -76.72
0.37334 | -76.9104
0.377 | -77.0023
0.38066 | -76.408
0.38432 | -76.1687
0.38798 | -75.9648
0.39164 | -76.8654
0.3953 | -76.8149
0.39896 | -76.0022
0.40262 | -76.9491
0.40628 | -77.8143
0.40994 | -77.2304
0.4136 | -77.6601
0.41726 | -78.0955
0.42092 | -76.5004
0.42458 | -76.9396
0.42824 | -77.1898
0.4319 | -77.5871
0.43556 | -78.5296
0.43922 | -78.6451
0.44288 | -77.4672
0.44654 | -77.8438
0.4502 | -79.09
0.45386 | -78.7956
0.45752 | -78.0483
0.46118 | -78.5099
0.46484 | -80.0778
0.4685 | -79.7166
0.47216 | -80.0119
0.47582 | -79.4188
0.47948 | -78.6648
0.48314 | -78.1302
0.4868 | -78.4616
0.49046 | -79.6216
0.49412 | -80.525
0.49778 | -80.0413
0.50144 | -79.2247
0.5051 | -78.2043
0.50876 | -78.0243
0.51242 | -79.0268
0.51608 | -79.9419
0.51974 | -79.8009
0.5234 | -80.4405
0.52706 | -79.9704
0.53072 | -79.6752
0.53438 | -79.3713
0.53804 | -79.1804
0.5417 | -79.5609
0.54536 | -80.1481
0.54902 | -80.6419
0.55268 | -80.5435
0.55634 | -80.2509
0.56 | -80.2082
0.56366 | -80.1072
0.56732 | -80.2355
0.57098 | -80.6449
0.57464 | -80.8851
0.5783 | -81.1675
0.58196 | -81.6887
0.58562 | -81.8426
0.58928 | -81.5064
0.59295 | -81.503
0.59661 | -81.7008
0.60027 | -81.8797
0.60393 | -81.9804
0.60759 | -81.678
0.61125 | -81.2505
0.61491 | -81.2912
0.61857 | -82.0293
0.62223 | -82.6774
0.62589 | -82.5781
0.62955 | -82.7454
0.63321 | -83.0921
0.63687 | -82.6192
0.64053 | -82.2144
0.64419 | -82.101
0.64785 | -81.7784
0.65151 | -81.5438
0.65517 | -81.9153
0.65883 | -82.5729
0.66249 | -83.145
0.66615 | -83.3416
0.66981 | -83.2074
0.67347 | -82.7019
0.67713 | -82.3263
0.68079 | -82.4926
0.68445 | -83.0172
0.68811 | -82.7929
0.69177 | -82.1357
0.69543 | -82.1188
0.69909 | -82.8103
0.70275 | -83.48
0.70641 | -83.641
0.71007 | -83.5445
0.71373 | -83.3698
0.71739 | -83.5133
0.72105 | -83.7712
0.72471 | -83.4634
0.72837 | -83.2549
0.73203 | -83.4579
0.73569 | -83.592
0.73935 | -83.6851
0.74301 | -84.1734
0.74667 | -84.6194
0.75033 | -84.2465
0.75399 | -83.6911
0.75765 | -83.6633
0.76131 | -83.7735
0.76497 | -83.6842
0.76863 | -83.8122
0.77229 | -84.296
0.77595 | -84.7379
0.77961 | -84.8014
0.78327 | -84.4252
0.78693 | -84.1245
0.79059 | -84.4801
0.79425 | -85.039
0.79791 | -85.0642
0.80157 | -84.8355
0.80523 | -84.4439
0.80889 | -83.7869
0.81255 | -83.3903
0.81621 | -83.6206
0.81988 | -84.5171
0.82354 | -85.4785
0.8272 | -86.097
0.83086 | -86.3918
0.83452 | -86.0879
0.83818 | -85.3194
0.84184 | -84.8645
0.8455 | -84.9959
0.84916 | -85.475
0.85282 | -85.6054
0.85648 | -85.1109
0.86014 | -84.8849
0.8638 | -85.1476
0.86746 | -85.1262
0.87112 | -85.0178
0.87478 | -84.9011
0.87844 | -84.6803
0.8821 | -84.5325
0.88576 | -84.9671
0.88942 | -85.7247
0.89308 | -86.5348
0.89674 | -86.9793
0.9004 | -86.6705
0.90406 | -85.958
0.90772 | -85.3355
0.91138 | -85.4956
0.91504 | -86.2362
0.9187 | -86.7416
0.92236 | -86.4852
0.92602 | -85.8734
0.92968 | -85.6752
0.93334 | -85.8476
0.937 | -85.7867
0.94066 | -85.5251
0.94432 | -85.5722
0.94798 | -85.9266
0.95164 | -86.3296
0.9553 | -86.512
0.95896 | -86.3321
0.96262 | -85.6482
0.96628 | -85.4854
0.96994 | -85.8451
0.9736 | -85.5684
0.97726 | -85.124
0.98092 | -85.2528
0.98458 | -85.9416
0.98824 | -86.2341
0.9919 | -85.9901
0.99556 | -86.178
0.99922 | -86.6702
1.0029 | -86.819
1.0065 | -86.7052
1.0102 | -86.9963
1.0139 | -87.1585
1.0175 | -86.7654
1.0212 | -86.489
1.0248 | -86.6948
1.0285 | -87.1213
1.0322 | -87.2641
1.0358 | -87.0666
1.0395 | -86.988
1.0431 | -87.2365
1.0468 | -87.8008
1.0505 | -88.1507
1.0541 | -87.5264
1.0578 | -87.1221
1.0614 | -86.951
1.0651 | -87.0174
1.0688 | -87.1601
1.0724 | -87.2406
1.0761 | -87.2677
1.0797 | -87.3021
1.0834 | -87.3781
1.0871 | -87.5044
1.0907 | -87.6368
1.0944 | -87.7749
1.098 | -87.9131
1.1017 | -87.9086
1.1054 | -87.7219
1.109 | -87.6045
1.1127 | -87.7277
1.1163 | -87.8482
1.12 | -87.7728
1.1237 | -87.6985
1.1273 | -87.7339
1.131 | -87.7893
1.1346 | -87.7762
1.1383 | -87.6253
1.142 | -87.3786
1.1456 | -87.308
1.1493 | -87.6408
1.1529 | -88.2494
1.1566 | -88.7941
1.1603 | -89.064
1.1639 | -88.9638
1.1676 | -88.5985
1.1712 | -88.247
1.1749 | -88.1279
1.1786 | -88.2523
1.1822 | -88.4174
1.1859 | -88.4363
1.1896 | -88.4135
1.1932 | -88.4635
1.1969 | -88.5022
1.2005 | -88.5059
1.2042 | -88.5622
1.2079 | -88.5415
1.2115 | -88.4638
1.2152 | -88.6708
1.2188 | -89.216
1.2225 | -89.6362
1.2262 | -89.5881
1.2298 | -89.2881
1.2335 | -88.9292
1.2371 | -88.5114
1.2408 | -88.2266
1.2445 | -88.2986
1.2481 | -88.7036
1.2518 | -89.144
1.2554 | -89.2631
1.2591 | -89.1485
1.2628 | -89.1702
1.2664 | -89.2618
1.2701 | -89.0894
1.2737 | -88.82
1.2774 | -88.7729
1.2811 | -88.9695
1.2847 | -89.2759
1.2884 | -89.6167
1.292 | -89.9663
1.2957 | -90.1115
1.2994 | -89.9194
1.303 | -89.7029
1.3067 | -89.7149
1.3103 | -89.8329
1.314 | -89.7663
1.3177 | -89.6491
1.3213 | -89.8426
1.325 | -90.2891
1.3286 | -90.6077
1.3323 | -90.6565
1.336 | -90.5338
1.3396 | -90.163
1.3433 | -89.5223
1.3469 | -88.9805
1.3506 | -88.8205
1.3543 | -88.9565
1.3579 | -89.159
1.3616 | -89.3084
1.3652 | -89.346
1.3689 | -89.2609
1.3726 | -89.211
1.3762 | -89.3568
1.3799 | -89.5945
1.3835 | -89.7016
1.3872 | -89.7137
1.3909 | -89.7827
1.3945 | -89.9041
1.3982 | -90.0436
1.4018 | -90.2208
1.4055 | -90.294
1.4092 | -90.139
1.4128 | -89.9582
1.4165 | -89.8758
1.4201 | -89.8094
1.4238 | -89.7453
1.4275 | -89.7532
1.4311 | -89.9351
1.4348 | -90.2115
1.4384 | -90.3118
1.4421 | -90.2286
1.4458 | -90.1683
1.4494 | -90.106
1.4531 | -89.8945
1.4567 | -89.6548
1.4604 | -89.6542
1.4641 | -89.9284
1.4677 | -90.248
1.4714 | -90.3629
1.475 | -90.3051
1.4787 | -90.27
1.4824 | -90.3638
1.486 | -90.5994
1.4897 | -90.8576
1.4933 | -90.8661
1.497 | -90.5943
1.5007 | -90.3579
1.5043 | -90.3754
1.508 | -90.6415
1.5116 | -90.9971
1.5153 | -91.1989
1.519 | -91.201
1.5226 | -91.2225
1.5263 | -91.3988
1.5299 | -91.5519
1.5336 | -91.4007
1.5373 | -90.9709
1.5409 | -90.5761
1.5446 | -90.4346
1.5482 | -90.5004
1.5519 | -90.7279
1.5556 | -91.1364
1.5592 | -91.5098
1.5629 | -91.6484
1.5665 | -91.676
1.5702 | -91.5988
1.5739 | -91.2882
1.5775 | -90.9097
1.5812 | -90.847
1.5848 | -90.889
1.5885 | -91.0653
1.5922 | -91.2022
1.5958 | -91.104
1.5995 | -90.8071
1.6031 | -90.5737
1.6068 | -90.584
1.6105 | -90.8044
1.6141 | -91.0821
1.6178 | -91.3091
1.6214 | -91.5485
1.6251 | -91.8321
1.6288 | -92.0733
1.6324 | -92.1775
1.6361 | -92.1792
1.6397 | -92.1509
1.6434 | -92.1194
1.6471 | -92.1242
1.6507 | -92.2145
1.6544 | -92.323
1.6581 | -92.3205
1.6617 | -92.1492
1.6654 | -91.8708
1.669 | -91.5998
1.6727 | -91.4414
1.6764 | -91.4308
1.68 | -91.5168
1.6837 | -91.5527
1.6873 | -91.4138
1.691 | -91.1451
1.6947 | -90.9383
1.6983 | -90.9506
1.702 | -91.1651
1.7056 | -91.4307
1.7093 | -91.6127
1.713 | -91.6782
1.7166 | -91.6793
1.7203 | -91.6936
1.7239 | -91.7398
1.7276 | -91.7306
1.7313 | -91.5737
1.7349 | -91.3362
1.7386 | -91.195
1.7422 | -91.2621
1.7459 | -91.4955
1.7496 | -91.7362
1.7532 | -91.7849
1.7569 | -91.6489
1.7605 | -91.531
1.7642 | -91.5933
1.7679 | -91.7589
1.7715 | -91.8273
1.7752 | -91.7303
1.7788 | -91.6064
1.7825 | -91.6105
1.7862 | -91.8028
1.7898 | -92.1159
1.7935 | -92.4039
1.7971 | -92.5682
1.8008 | -92.6369
1.8045 | -92.6652
1.8081 | -92.6758
1.8118 | -92.6361
1.8154 | -92.4677
1.8191 | -92.1969
1.8228 | -91.9054
1.8264 | -91.6803
1.8301 | -91.5383
1.8337 | -91.4671
1.8374 | -91.4789
1.8411 | -91.6064
1.8447 | -91.854
1.8484 | -92.1488
1.852 | -92.3746
1.8557 | -92.4748
1.8594 | -92.4395
1.863 | -92.3499
1.8667 | -92.3259
1.8703 | -92.4325
1.874 | -92.6419
1.8777 | -92.8474
1.8813 | -92.9511
1.885 | -92.9658
1.8886 | -92.951
1.8923 | -92.903
1.896 | -92.7728
1.8996 | -92.5795
1.9033 | -92.3812
1.9069 | -92.2249
1.9106 | -92.1231
1.9143 | -92.0583
1.9179 | -92.013
1.9216 | -92.0127
1.9252 | -92.0782
1.9289 | -92.1533
1.9326 | -92.1765
1.9362 | -92.1911
1.9399 | -92.266
1.9435 | -92.3801
1.9472 | -92.4677
1.9509 | -92.4716
1.9545 | -92.3978
1.9582 | -92.3031
1.9618 | -92.2946
1.9655 | -92.3969
1.9692 | -92.5174
1.9728 | -92.5341
1.9765 | -92.4332
1.9801 | -92.28
1.9838 | -92.1529
1.9875 | -92.1093
1.9911 | -92.2168
1.9948 | -92.4596
1.9984 | -92.7445
2.0021 | -92.9343
2.0058 | -92.9692
2.0094 | -92.8898
2.0131 | -92.8028
2.0167 | -92.787
2.0204 | -92.8632
2.0241 | -93.0123
2.0277 | -93.2041
2.0314 | -93.3869
2.035 | -93.5032
2.0387 | -93.5088
2.0424 | -93.4374
2.046 | -93.3149
2.0497 | -93.1501
2.0533 | -92.9458
2.057 | -92.7414
2.0607 | -92.5923
2.0643 | -92.5566
2.068 | -92.6725
2.0716 | -92.9199
2.0753 | -93.2266
2.079 | -93.5052
2.0826 | -93.6632
2.0863 | -93.6955
2.0899 | -93.6596
2.0936 | -93.6103
2.0973 | -93.5184
2.1009 | -93.3521
2.1046 | -93.1528
2.1082 | -92.9662
2.1119 | -92.8514
2.1156 | -92.765
2.1192 | -92.7569
2.1229 | -92.8233
2.1265 | -92.9457
2.1302 | -93.0689
2.1339 | -93.1243
2.1375 | -93.0997
2.1412 | -93.041
2.1449 | -93.0229
2.1485 | -93.092
2.1522 | -93.2594
2.1558 | -93.4661
2.1595 | -93.637
2.1632 | -93.7439
2.1668 | -93.7646
2.1705 | -93.7046
2.1741 | -93.6188
2.1778 | -93.5538
2.1815 | -93.5334
2.1851 | -93.5488
2.1888 | -93.5749
2.1924 | -93.6093
2.1961 | -93.6588
2.1998 | -93.696
2.2034 | -93.6764
2.2071 | -93.603
2.2107 | -93.4926
2.2144 | -93.3828
2.2181 | -93.3187
2.2217 | -93.3257
2.2254 | -93.3819
2.229 | -93.4447
2.2327 | -93.4919
2.2364 | -93.5367
2.24 | -93.6116
2.2437 | -93.6959
2.2473 | -93.7524
2.251 | -93.7353
2.2547 | -93.6449
2.2583 | -93.5551
2.262 | -93.5488
2.2656 | -93.6633
2.2693 | -93.8662
2.273 | -94.0798
2.2766 | -94.2303
2.2803 | -94.2717
2.2839 | -94.2353
2.2876 | -94.1812
2.2913 | -94.1785
2.2949 | -94.2538
2.2986 | -94.3783
2.3022 | -94.4843
2.3059 | -94.5116
2.3096 | -94.4199
2.3132 | -94.2645
2.3169 | -94.1019
2.3205 | -93.9616
2.3242 | -93.8648
2.3279 | -93.8238
2.3315 | -93.8448
2.3352 | -93.8715
2.3388 | -93.8317
2.3425 | -93.7102
2.3462 | -93.5771
2.3498 | -93.5371
2.3535 | -93.6021
2.3571 | -93.7311
2.3608 | -93.8656
2.3645 | -93.9686
2.3681 | -94.0482
2.3718 | -94.0911
2.3754 | -94.1025
2.3791 | -94.0976
2.3828 | -94.0756
2.3864 | -94.0265
2.3901 | -93.96
2.3937 | -93.9055
2.3974 | -93.8905
2.4011 | -93.925
2.4047 | -93.9992
2.4084 | -94.1113
2.412 | -94.2691
2.4157 | -94.4489
2.4194 | -94.6061
2.423 | -94.6858
2.4267 | -94.6383
2.4303 | -94.464
2.434 | -94.2368
2.4377 | -94.0524
2.4413 | -93.993
2.445 | -94.0589
2.4486 | -94.1495
2.4523 | -94.1867
2.456 | -94.1719
2.4596 | -94.1736
2.4633 | -94.2512
2.4669 | -94.4068
2.4706 | -94.6125
2.4743 | -94.8159
2.4779 | -94.9757
2.4816 | -95.0688
2.4852 | -95.1061
2.4889 | -95.1158
2.4926 | -95.0915
2.4962 | -95.0643
2.4999 | -95.016
2.5035 | -94.9416
2.5072 | -94.8555
2.5109 | -94.773
2.5145 | -94.7211
2.5182 | -94.7183
2.5218 | -94.7642
2.5255 | -94.8325
2.5292 | -94.8696
2.5328 | -94.8452
2.5365 | -94.761
2.5401 | -94.6874
2.5438 | -94.6607
2.5475 | -94.7
2.5511 | -94.8003
2.5548 | -94.9198
2.5584 | -94.9819
2.5621 | -94.9414
2.5658 | -94.8186
2.5694 | -94.6903
2.5731 | -94.6114
2.5767 | -94.6126
2.5804 | -94.7083
2.5841 | -94.8872
2.5877 | -95.1032
2.5914 | -95.3035
2.5951 | -95.4424
2.5987 | -95.4898
2.6024 | -95.4342
2.606 | -95.2613
2.6097 | -94.9894
2.6134 | -94.7239
2.617 | -94.5754
2.6207 | -94.5831
2.6243 | -94.732
2.628 | -94.9342
2.6317 | -95.0686
2.6353 | -95.0787
2.639 | -95.0214
2.6426 | -95.0188
2.6463 | -95.028
2.65 | -95.0452
2.6536 | -95.0378
2.6573 | -94.9739
2.6609 | -94.8431
2.6646 | -94.6793
2.6683 | -94.5169
2.6719 | -94.3973
2.6756 | -94.3459
2.6792 | -94.3595
2.6829 | -94.4178
2.6866 | -94.517
2.6902 | -94.6263
2.6939 | -94.7275
2.6975 | -94.8155
2.7012 | -94.881
2.7049 | -94.926
2.7085 | -94.9844
2.7122 | -95.0755
2.7158 | -95.2044
2.7195 | -95.3603
2.7232 | -95.513
2.7268 | -95.6232
2.7305 | -95.6529
2.7341 | -95.5835
2.7378 | -95.4606
2.7415 | -95.311
2.7451 | -95.1842
2.7488 | -95.123
2.7524 | -95.1171
2.7561 | -95.146
2.7598 | -95.198
2.7634 | -95.2578
2.7671 | -95.3075
2.7707 | -95.3521
2.7744 | -95.397
2.7781 | -95.4617
2.7817 | -95.5818
2.7854 | -95.7572
2.789 | -95.9822
2.7927 | -96.2112
2.7964 | -96.3655
2.8 | -96.4218
2.8037 | -96.3742
2.8073 | -96.2399
2.811 | -96.0534
2.8147 | -95.8683
2.8183 | -95.7182
2.822 | -95.6382
2.8256 | -95.6357
2.8293 | -95.6918
2.833 | -95.7907
2.8366 | -95.8874
2.8403 | -95.9324
2.8439 | -95.9194
2.8476 | -95.8572
2.8513 | -95.7769
2.8549 | -95.7223
2.8586 | -95.6894
2.8622 | -95.6491
2.8659 | -95.5985
2.8696 | -95.5462
2.8732 | -95.5098
2.8769 | -95.4884
2.8805 | -95.466
2.8842 | -95.4409
2.8879 | -95.4121
2.8915 | -95.3664
2.8952 | -95.2964
2.8988 | -95.2047
2.9025 | -95.111
2.9062 | -95.0611
2.9098 | -95.1028
2.9135 | -95.2522
2.9171 | -95.4795
2.9208 | -95.7267
2.9245 | -95.9351
2.9281 | -96.0479
2.9318 | -96.0434
2.9354 | -95.9448
2.9391 | -95.8056
2.9428 | -95.6831
2.9464 | -95.6257
2.9501 | -95.6196
2.9537 | -95.6585
2.9574 | -95.6996
2.9611 | -95.692
2.9647 | -95.6458
2.9684 | -95.5842
2.972 | -95.5398
2.9757 | -95.5428
2.9794 | -95.6101
2.983 | -95.7345
2.9867 | -95.8949
2.9903 | -96.0586
2.994 | -96.203
2.9977 | -96.3262
3.0013 | -96.3909
3.005 | -96.3993
3.0086 | -96.3755
3.0123 | -96.3209
3.016 | -96.2541
3.0196 | -96.1997
3.0233 | -96.1606
3.0269 | -96.1217
3.0306 | -96.077
3.0343 | -96.0293
3.0379 | -95.9763
3.0416 | -95.9023
3.0452 | -95.7963
3.0489 | -95.6857
3.0526 | -95.6117
3.0562 | -95.6063
3.0599 | -95.6762
3.0635 | -95.7961
3.0672 | -95.902
3.0709 | -95.941
3.0745 | -95.8973
3.0782 | -95.8181
3.0819 | -95.7735
3.0855 | -95.8159
3.0892 | -95.9562
3.0928 | -96.1654
3.0965 | -96.366
3.1002 | -96.5083
3.1038 | -96.5865
3.1075 | -96.6197
3.1111 | -96.6488
3.1148 | -96.6961
3.1185 | -96.7551
3.1221 | -96.8063
3.1258 | -96.8256
3.1294 | -96.8143
3.1331 | -96.7948
3.1368 | -96.7866
3.1404 | -96.7867
3.1441 | -96.8105
3.1477 | -96.8653
3.1514 | -96.9673
3.1551 | -97.1378
3.1587 | -97.3551
3.1624 | -97.5786
3.166 | -97.6141
3.1697 | -97.6229
3.1734 | -97.5336
3.177 | -97.3705
3.1807 | -97.1682
3.1843 | -96.9666
3.188 | -96.8123
3.1917 | -96.7408
3.1953 | -96.7638
3.199 | -96.876
3.2026 | -97.0479
3.2063 | -97.2352
3.21 | -97.3898
3.2136 | -97.4831
3.2173 | -97.4935
3.2209 | -97.4286
3.2246 | -97.3084
3.2283 | -97.1716
3.2319 | -97.0581
3.2356 | -97.0064
3.2392 | -97.0162
3.2429 | -97.0648
3.2466 | -97.1289
3.2502 | -97.1868
3.2539 | -97.2214
3.2575 | -97.2366
3.2612 | -97.219
3.2649 | -97.179
3.2685 | -97.1143
3.2722 | -97.038
3.2758 | -96.9693
3.2795 | -96.934
3.2832 | -96.9458
3.2868 | -97.0022
3.2905 | -97.0829
3.2941 | -97.179
3.2978 | -97.2641
3.3015 | -97.3102
3.3051 | -97.3066
3.3088 | -97.259
3.3124 | -97.1728
3.3161 | -97.0641
3.3198 | -96.9539
3.3234 | -96.8746
3.3271 | -96.8508
3.3307 | -96.8944
3.3344 | -96.9936
3.3381 | -97.1222
3.3417 | -97.2598
3.3454 | -97.3919
3.349 | -97.4952
3.3527 | -97.5656
3.3564 | -97.6204
3.36 | -97.6618
3.3637 | -97.6951
3.3673 | -97.7338
3.371 | -97.7772
3.3747 | -97.8225
3.3783 | -97.8479
3.382 | -97.851
3.3856 | -97.8033
3.3893 | -97.7012
3.393 | -97.5522
3.3966 | -97.364
3.4003 | -97.1524
3.4039 | -96.9549
3.4076 | -96.7964
3.4113 | -96.6901
3.4149 | -96.6615
3.4186 | -96.7087
3.4222 | -96.8204
3.4259 | -96.97
3.4296 | -97.1131
3.4332 | -97.2245
3.4369 | -97.297
3.4405 | -97.3559
3.4442 | -97.4166
3.4479 | -97.498
3.4515 | -97.6014
3.4552 | -97.7131
3.4588 | -97.8171
3.4625 | -97.8949
3.4662 | -97.9233
3.4698 | -97.9019
3.4735 | -97.8521
3.4771 | -97.7891
3.4808 | -97.7311
3.4845 | -97.6926
3.4881 | -97.681
3.4918 | -97.6887
3.4954 | -97.7108
3.4991 | -97.7411
3.5028 | -97.7677
3.5064 | -97.7948
3.5101 | -97.8383
3.5137 | -97.8987
3.5174 | -97.9762
3.5211 | -98.0677
3.5247 | -98.1385
3.5284 | -98.1472
3.532 | -98.0868
3.5357 | -97.9646
3.5394 | -97.795
3.543 | -97.6045
3.5467 | -97.4232
3.5503 | -97.2775
3.554 | -97.1831
3.5577 | -97.1342
3.5613 | -97.1361
3.565 | -97.1724
3.5687 | -97.2275
3.5723 | -97.2858
3.576 | -97.3338
3.5796 | -97.3756
3.5833 | -97.4235
3.587 | -97.4814
3.5906 | -97.5522
3.5943 | -97.6343
3.5979 | -97.7058
3.6016 | -97.7506
3.6053 | -97.7639
3.6089 | -97.7507
3.6126 | -97.7177
3.6162 | -97.6737
3.6199 | -97.6424
3.6236 | -97.6498
3.6272 | -97.7113
3.6309 | -97.8265
3.6345 | -97.986
3.6382 | -98.1573
3.6419 | -98.3017
3.6455 | -98.3913
3.6492 | -98.4099
3.6528 | -98.359
3.6565 | -98.2511
3.6602 | -98.126
3.6638 | -98.0103
3.6675 | -97.9309
3.6711 | -97.8894
3.6748 | -97.8962
3.6785 | -97.9364
3.6821 | -98.0027
3.6858 | -98.0885
3.6894 | -98.1685
3.6931 | -98.2291
3.6968 | -98.2947
3.7004 | -98.3455
3.7041 | -98.3855
3.7077 | -98.4191
3.7114 | -98.4449
3.7151 | -98.4652
3.7187 | -98.4729
3.7224 | -98.4797
3.726 | -98.4763
3.7297 | -98.4676
3.7334 | -98.451
3.737 | -98.418
3.7407 | -98.3675
3.7443 | -98.3024
3.748 | -98.237
3.7517 | -98.1742
3.7553 | -98.1265
3.759 | -98.099
3.7626 | -98.0886
3.7663 | -98.0927
3.77 | -98.1131
3.7736 | -98.1447
3.7773 | -98.1605
3.7809 | -98.1572
3.7846 | -98.1358
3.7883 | -98.0908
3.7919 | -98.0208
3.7956 | -97.9376
3.7992 | -97.8708
3.8029 | -97.8363
3.8066 | -97.8548
3.8102 | -97.9312
3.8139 | -98.0561
3.8175 | -98.2115
3.8212 | -98.3736
3.8249 | -98.5074
3.8285 | -98.5833
3.8322 | -98.5932
3.8358 | -98.5538
3.8395 | -98.4842
3.8432 | -98.3973
3.8468 | -98.3137
3.8505 | -98.2606
3.8541 | -98.2523
3.8578 | -98.2913
3.8615 | -98.3748
3.8651 | -98.4986
3.8688 | -98.6476
3.8724 | -98.8103
3.8761 | -98.9746
3.8798 | -99.1174
3.8834 | -99.2151
3.8871 | -99.2523
3.8907 | -99.233
3.8944 | -99.1624
3.8981 | -99.0487
3.9017 | -98.922
3.9054 | -98.7977
3.909 | -98.69
3.9127 | -98.5987
3.9164 | -98.5172
3.92 | -98.451
3.9237 | -98.3874
3.9273 | -98.3286
3.931 | -98.2795
3.9347 | -98.2355
3.9383 | -98.1987
3.942 | -98.1718
3.9456 | -98.1612
3.9493 | -98.1751
3.953 | -98.2178
3.9566 | -98.2912
3.9603 | -98.3987
3.9639 | -98.5245
3.9676 | -98.6558
3.9713 | -98.7702
3.9749 | -98.8648
3.9786 | -98.9292
3.9822 | -98.955
3.9859 | -98.9605
3.9896 | -98.9377
3.9932 | -98.9061
3.9969 | -98.8844
4.0005 | -98.88
4.0042 | -98.8928
4.0079 | -98.9132
4.0115 | -98.9445
4.0152 | -98.982
4.0188 | -99.0306
4.0225 | -99.0744
4.0262 | -99.1039
4.0298 | -99.1226
4.0335 | -99.1182
4.0372 | -99.0862
4.0408 | -99.0265
4.0445 | -98.9552
4.0481 | -98.8854
4.0518 | -98.8298
4.0555 | -98.7937
4.0591 | -98.7602
4.0628 | -98.7178
4.0664 | -98.6682
4.0701 | -98.6064
4.0738 | -98.5355
4.0774 | -98.4802
4.0811 | -98.454
4.0847 | -98.4698
4.0884 | -98.5196
4.0921 | -98.5829
4.0957 | -98.656
4.0994 | -98.7308
4.103 | -98.8078
4.1067 | -98.8785
4.1104 | -98.9363
4.114 | -98.9777
4.1177 | -99.0038
4.1213 | -99.0236
4.125 | -99.0413
4.1287 | -99.055
4.1323 | -99.0785
4.136 | -99.1046
4.1396 | -99.1296
4.1433 | -99.158
4.147 | -99.1861
4.1506 | -99.209
4.1543 | -99.232
4.1579 | -99.261
4.1616 | -99.282
4.1653 | -99.2979
4.1689 | -99.307
4.1726 | -99.3116
4.1762 | -99.3139
4.1799 | -99.2839
4.1836 | -99.2306
4.1872 | -99.1516
4.1909 | -99.062
4.1945 | -98.9832
4.1982 | -98.9375
4.2019 | -98.9325
4.2055 | -98.9779
4.2092 | -99.0615
4.2128 | -99.1751
4.2165 | -99.2962
4.2202 | -99.3926
4.2238 | -99.401
4.2275 | -99.4393
4.2311 | -99.4449
4.2348 | -99.4258
4.2385 | -99.3863
4.2421 | -99.3397
4.2458 | -99.3029
4.2494 | -99.2766
4.2531 | -99.2702
4.2568 | -99.287
4.2604 | -99.3182
4.2641 | -99.3651
4.2677 | -99.4245
4.2714 | -99.4751
4.2751 | -99.5184
4.2787 | -99.5408
4.2824 | -99.5254
4.286 | -99.4854
4.2897 | -99.4171
4.2934 | -99.3206
4.297 | -99.2184
4.3007 | -99.1333
4.3043 | -99.0558
4.308 | -98.9875
4.3117 | -98.9307
4.3153 | -98.9003
4.319 | -98.8995
4.3226 | -98.9335
4.3263 | -98.996
4.33 | -99.0843
4.3336 | -99.1865
4.3373 | -99.287
4.3409 | -99.3729
4.3446 | -99.4308
4.3483 | -99.4603
4.3519 | -99.4719
4.3556 | -99.4617
4.3592 | -99.4494
4.3629 | -99.4333
4.3666 | -99.4146
4.3702 | -99.3908
4.3739 | -99.3656
4.3775 | -99.3491
4.3812 | -99.333
4.3849 | -99.3157
4.3885 | -99.3015
4.3922 | -99.2934
4.3958 | -99.2955
4.3995 | -99.302
4.4032 | -99.3051
4.4068 | -99.3062
4.4105 | -99.3151
4.4141 | -99.3266
4.4178 | -99.3451
4.4215 | -99.3843
4.4251 | -99.4418
4.4288 | -99.5078
4.4324 | -99.5898
4.4361 | -99.6817
4.4398 | -99.7847
4.4434 | -99.8843
4.4471 | -99.9602
4.4507 | -100.027
4.4544 | -100.0775
4.4581 | -100.0941
4.4617 | -100.0671
4.4654 | -100.0091
4.469 | -99.9373
4.4727 | -99.8537
4.4764 | -99.7631
4.48 | -99.6757
4.4837 | -99.6038
4.4874 | -99.5527
4.491 | -99.5184
4.4947 | -99.5071
4.4983 | -99.5126
4.502 | -99.5274
4.5057 | -99.548
4.5093 | -99.5775
4.513 | -99.618
4.5166 | -99.6709
4.5203 | -99.7341
4.524 | -99.8122
4.5276 | -99.9115
4.5313 | -100.0171
4.5349 | -100.1146
4.5386 | -100.191
4.5423 | -100.2488
4.5459 | -100.283
4.5496 | -100.2804
4.5532 | -100.2511
4.5569 | -100.1927
4.5606 | -100.1157
4.5642 | -100.0421
4.5679 | -99.9858
4.5715 | -99.9508
4.5752 | -99.9383
4.5789 | -99.9526
4.5825 | -99.9761
4.5862 | -99.9931
4.5898 | -100.0064
4.5935 | -100.0111
4.5972 | -100.0093
4.6008 | -100.0052
4.6045 | -100.0015
4.6081 | -100.0076
4.6118 | -100.0235
4.6155 | -100.0471
4.6191 | -100.0864
4.6228 | -100.1268
4.6264 | -100.16
4.6301 | -100.1745
4.6338 | -100.1555
4.6374 | -100.1087
4.6411 | -100.0429
4.6447 | -99.9713
4.6484 | -99.8989
4.6521 | -99.8527
4.6557 | -99.8367
4.6594 | -99.8393
4.663 | -99.8601
4.6667 | -99.891
4.6704 | -99.9339
4.674 | -99.9758
4.6777 | -99.9985
4.6813 | -100.0052
4.685 | -99.9988
4.6887 | -99.9745
4.6923 | -99.9447
4.696 | -99.9183
4.6996 | -99.9015
4.7033 | -99.8994
4.707 | -99.9101
4.7106 | -99.9425
4.7143 | -99.9933
4.7179 | -100.0618
4.7216 | -100.149
4.7253 | -100.2391
4.7289 | -100.3163
4.7326 | -100.3724
4.7362 | -100.407
4.7399 | -100.4079
4.7436 | -100.3816
4.7472 | -100.334
4.7509 | -100.2439
4.7545 | -100.1801
4.7582 | -100.1102
4.7619 | -100.0421
4.7655 | -99.9817
4.7692 | -99.9311
4.7728 | -99.893
4.7765 | -99.8777
4.7802 | -99.8907
4.7838 | -99.9284
4.7875 | -99.9972
4.7911 | -100.0942
4.7948 | -100.2119
4.7985 | -100.3483
4.8021 | -100.4961
4.8058 | -100.6438
4.8094 | -100.7673
4.8131 | -100.8715
4.8168 | -100.9493
4.8204 | -100.9967
4.8241 | -101.0091
4.8277 | -100.9924
4.8314 | -100.9455
4.8351 | -100.8723
4.8387 | -100.7857
4.8424 | -100.6897
4.846 | -100.5873
4.8497 | -100.4837
4.8534 | -100.3804
4.857 | -100.2922
4.8607 | -100.2238
4.8643 | -100.1816
4.868 | -100.1626
4.8717 | -100.1696
4.8753 | -100.2055
4.879 | -100.2574
4.8826 | -100.314
4.8863 | -100.3695
4.89 | -100.4134
4.8936 | -100.4505
4.8973 | -100.4848
4.9009 | -100.5082
4.9046 | -100.5299
4.9083 | -100.555
4.9119 | -100.5877
4.9156 | -100.6236
4.9192 | -100.6604
4.9229 | -100.6958
4.9266 | -100.7191
4.9302 | -100.7259
4.9339 | -100.7045
4.9375 | -100.6638
4.9412 | -100.6075
4.9449 | -100.5387
4.9485 | -100.4592
4.9522 | -100.385
4.9558 | -100.3248
4.9595 | -100.2786
4.9632 | -100.2513
4.9668 | -100.2446
4.9705 | -100.2536
4.9741 | -100.2717
4.9778 | -100.3078
4.9815 | -100.3501
4.9851 | -100.3992
4.9888 | -100.4477
4.9925 | -100.4927
4.9961 | -100.5228
4.9998 | -100.545
5.0034 | -100.5558
5.0071 | -100.5533
5.0108 | -100.5555
5.0144 | -100.5607
5.0181 | -100.5731
5.0217 | -100.5961
5.0254 | -100.6259
5.0291 | -100.6617
5.0327 | -100.6976
5.0364 | -100.7339
5.04 | -100.7628
5.0437 | -100.7734
5.0474 | -100.7727
5.051 | -100.7697
5.0547 | -100.7506
5.0583 | -100.7242
5.062 | -100.702
5.0657 | -100.6808
5.0693 | -100.6582
5.073 | -100.65
5.0766 | -100.6574
5.0803 | -100.6829
5.084 | -100.7295
5.0876 | -100.7905
5.0913 | -100.8671
5.0949 | -100.9577
5.0986 | -101.0574
5.1023 | -101.1412
5.1059 | -101.202
5.1096 | -101.2399
5.1132 | -101.2428
5.1169 | -101.2248
5.1206 | -101.1957
5.1242 | -101.1558
5.1279 | -101.1181
5.1315 | -101.0929
5.1352 | -101.0631
5.1389 | -101.0308
5.1425 | -101.0058
5.1462 | -100.9758
5.1498 | -100.9278
5.1535 | -100.863
5.1572 | -100.7928
5.1608 | -100.7087
5.1645 | -100.6233
5.1681 | -100.5402
5.1718 | -100.4731
5.1755 | -100.4216
5.1791 | -100.3991
5.1828 | -100.3988
5.1864 | -100.4175
5.1901 | -100.4571
5.1938 | -100.5006
5.1974 | -100.549
5.2011 | -100.5965
5.2047 | -100.632
5.2084 | -100.6591
5.2121 | -100.6868
5.2157 | -100.7031
5.2194 | -100.7202
5.223 | -100.7466
5.2267 | -100.7796
5.2304 | -100.8258
5.234 | -100.8896
5.2377 | -100.9777
5.2413 | -101.068
5.245 | -101.1587
5.2487 | -101.2426
5.2523 | -101.3193
5.256 | -101.3733
5.2596 | -101.4101
5.2633 | -101.4396
5.267 | -101.4499
5.2706 | -101.4476
5.2743 | -101.4326
5.2779 | -101.4113
5.2816 | -101.392
5.2853 | -101.3724
5.2889 | -101.3486
5.2926 | -101.3221
5.2962 | -101.2925
5.2999 | -101.2431
5.3036 | -101.1916
5.3072 | -101.1466
5.3109 | -101.093
5.3145 | -101.0317
5.3182 | -100.9614
5.3219 | -100.8844
5.3255 | -100.8127
5.3292 | -100.7385
5.3328 | -100.6725
5.3365 | -100.6294
5.3402 | -100.6042
5.3438 | -100.5933
5.3475 | -100.6004
5.3511 | -100.6266
5.3548 | -100.6594
5.3585 | -100.6938
5.3621 | -100.7238
5.3658 | -100.7488
5.3694 | -100.763
5.3731 | -100.7707
5.3768 | -100.7746
5.3804 | -100.773
5.3841 | -100.7831
5.3877 | -100.789
5.3914 | -100.7859
5.3951 | -100.7882
5.3987 | -100.7955
5.4024 | -100.8091
5.406 | -100.8278
5.4097 | -100.8399
5.4134 | -100.8571
5.417 | -100.8805
5.4207 | -100.901
5.4243 | -100.9297
5.428 | -100.9707
5.4317 | -101.0219
5.4353 | -101.0769
5.439 | -101.1338
5.4427 | -101.1964
5.4463 | -101.264
5.45 | -101.3362
5.4536 | -101.4061
5.4573 | -101.4658
5.461 | -101.5144
5.4646 | -101.557
5.4683 | -101.5794
5.4719 | -101.5909
5.4756 | -101.5883
5.4793 | -101.5725
5.4829 | -101.5466
5.4866 | -101.511
5.4902 | -101.4805
5.4939 | -101.4574
5.4976 | -101.4407
5.5012 | -101.44
5.5049 | -101.4606
5.5085 | -101.496
5.5122 | -101.5488
5.5159 | -101.6077
5.5195 | -101.6663
5.5232 | -101.7253
5.5268 | -101.7652
5.5305 | -101.7856
5.5342 | -101.7839
5.5378 | -101.7586
5.5415 | -101.7194
5.5451 | -101.6647
5.5488 | -101.6017
5.5525 | -101.5366
5.5561 | -101.473
5.5598 | -101.416
5.5634 | -101.3602
5.5671 | -101.3148
5.5708 | -101.2823
5.5744 | -101.2511
5.5781 | -101.2233
5.5817 | -101.207
5.5854 | -101.1987
5.5891 | -101.1997
5.5927 | -101.2039
5.5964 | -101.1997
5.6 | -101.2008
5.6037 | -101.1985
5.6074 | -101.1892
5.611 | -101.179
5.6147 | -101.1649
5.6183 | -101.1518
5.622 | -101.1477
5.6257 | -101.1449
5.6293 | -101.1462
5.633 | -101.1527
5.6366 | -101.1589
5.6403 | -101.1647
5.644 | -101.169
5.6476 | -101.1678
5.6513 | -101.1604
5.6549 | -101.153
5.6586 | -101.15
5.6623 | -101.1511
5.6659 | -101.1508
5.6696 | -101.153
5.6732 | -101.1556
5.6769 | -101.1645
5.6806 | -101.1761
5.6842 | -101.183
5.6879 | -101.1913
5.6915 | -101.1978
5.6952 | -101.2044
5.6989 | -101.2148
5.7025 | -101.2146
5.7062 | -101.2027
5.7098 | -101.1811
5.7135 | -101.1517
5.7172 | -101.1121
5.7208 | -101.0737
5.7245 | -101.0384
5.7281 | -101.0155
5.7318 | -101.0089
5.7355 | -101.0186
5.7391 | -101.0453
5.7428 | -101.0822
5.7464 | -101.1281
5.7501 | -101.1785
5.7538 | -101.2343
5.7574 | -101.2901
5.7611 | -101.3355
5.7647 | -101.3649
5.7684 | -101.3757
5.7721 | -101.363
5.7757 | -101.3402
5.7794 | -101.3017
5.783 | -101.249
5.7867 | -101.1995
5.7904 | -101.1621
5.794 | -101.1376
5.7977 | -101.1296
5.8013 | -101.1331
5.805 | -101.1805
5.8087 | -101.2148
5.8123 | -101.2556
5.816 | -101.2922
5.8196 | -101.3288
5.8233 | -101.3671
5.827 | -101.3996
5.8306 | -101.4317
5.8343 | -101.4579
5.8379 | -101.479
5.8416 | -101.4862
5.8453 | -101.4877
5.8489 | -101.4918
5.8526 | -101.4907
5.8562 | -101.4797
5.8599 | -101.4624
5.8636 | -101.4296
5.8672 | -101.3954
5.8709 | -101.3576
5.8745 | -101.3187
5.8782 | -101.2771
5.8819 | -101.2371
5.8855 | -101.2019
5.8892 | -101.1755
5.8928 | -101.1475
5.8965 | -101.1188
5.9002 | -101.1002
5.9038 | -101.09
5.9075 | -101.0902
5.9112 | -101.0965
5.9148 | -101.1048
5.9185 | -101.1168
5.9221 | -101.1314
5.9258 | -101.1547
5.9295 | -101.1835
5.9331 | -101.2201
5.9368 | -101.2642
5.9404 | -101.3096
5.9441 | -101.3701
5.9478 | -101.443
5.9514 | -101.5198
5.9551 | -101.6052
5.9587 | -101.6891
5.9624 | -101.7708
5.9661 | -101.8427
5.9697 | -101.9049
5.9734 | -101.9521
5.977 | -101.9781
5.9807 | -101.9879
5.9844 | -101.9788
5.988 | -101.9492
5.9917 | -101.9077
5.9953 | -101.8685
5.999 | -101.8293
6.0027 | -101.7869
6.0063 | -101.7569
6.01 | -101.7477
6.0136 | -101.7509
6.0173 | -101.7657
6.021 | -101.7964
6.0246 | -101.8367
6.0283 | -101.8728
6.0319 | -101.9087
6.0356 | -101.9335
6.0393 | -101.9462
6.0429 | -101.9564
6.0466 | -101.9567
6.0502 | -101.9481
6.0539 | -101.9314
6.0576 | -101.9113
6.0612 | -101.8922
6.0649 | -101.8765
6.0685 | -101.8698
6.0722 | -101.8678
6.0759 | -101.8686
6.0795 | -101.8762
6.0832 | -101.8845
6.0868 | -101.892
6.0905 | -101.9075
6.0942 | -101.9368
6.0978 | -101.9627
6.1015 | -101.985
6.1051 | -102.0059
6.1088 | -102.0213
6.1125 | -102.0338
6.1161 | -102.0483
6.1198 | -102.0616
6.1234 | -102.0658
6.1271 | -102.0673
6.1308 | -102.0663
6.1344 | -102.0612
6.1381 | -102.0611
6.1417 | -102.0622
6.1454 | -102.0705
6.1491 | -102.0831
6.1527 | -102.0891
6.1564 | -102.0913
6.16 | -102.0888
6.1637 | -102.0856
6.1674 | -102.0797
6.171 | -102.0617
6.1747 | -102.0319
6.1783 | -102.0024
6.182 | -101.9753
6.1857 | -101.9455
6.1893 | -101.9212
6.193 | -101.9012
6.1966 | -101.8846
6.2003 | -101.8684
6.204 | -101.8555
6.2076 | -101.8445
6.2113 | -101.8388
6.2149 | -101.8412
6.2186 | -101.8494
6.2223 | -101.866
6.2259 | -101.8865
6.2296 | -101.9059
6.2332 | -101.9351
6.2369 | -101.9649
6.2406 | -101.9916
6.2442 | -102.0222
6.2479 | -102.0515
6.2515 | -102.0675
6.2552 | -102.0723
6.2589 | -102.0667
6.2625 | -102.0526
6.2662 | -102.0293
6.2698 | -101.9858
6.2735 | -101.9263
6.2772 | -101.87
6.2808 | -101.8126
6.2845 | -101.7543
6.2881 | -101.7021
6.2918 | -101.6547
6.2955 | -101.6146
6.2991 | -101.5903
6.3028 | -101.5844
6.3064 | -101.5957
6.3101 | -101.6259
6.3138 | -101.6779
6.3174 | -101.7517
6.3211 | -101.8454
6.3247 | -101.9488
6.3284 | -102.0568
6.3321 | -102.17
6.3357 | -102.2641
6.3394 | -102.3572
6.343 | -102.4401
6.3467 | -102.5159
6.3504 | -102.5753
6.354 | -102.6176
6.3577 | -102.643
6.3613 | -102.652
6.365 | -102.649
6.3687 | -102.6314
6.3723 | -102.5909
6.376 | -102.5429
6.3796 | -102.4951
6.3833 | -102.4405
6.387 | -102.3838
6.3906 | -102.3329
6.3943 | -102.2832
6.3979 | -102.2414
6.4016 | -102.2061
6.4053 | -102.1794
6.4089 | -102.1593
6.4126 | -102.1501
6.4163 | -102.1472
6.4199 | -102.148
6.4236 | -102.1524
6.4272 | -102.1518
6.4309 | -102.1453
6.4346 | -102.1513
6.4382 | -102.1555
6.4419 | -102.1514
6.4455 | -102.1469
6.4492 | -102.1514
6.4529 | -102.1626
6.4565 | -102.1715
6.4602 | -102.1745
6.4638 | -102.1778
6.4675 | -102.1845
6.4712 | -102.1901
6.4748 | -102.189
6.4785 | -102.1859
6.4821 | -102.1881
6.4858 | -102.1927
6.4895 | -102.1937
6.4931 | -102.2009
6.4968 | -102.2123
6.5004 | -102.2277
6.5041 | -102.2464
6.5078 | -102.265
6.5114 | -102.2842
6.5151 | -102.3059
6.5187 | -102.3309
6.5224 | -102.3533
6.5261 | -102.3808
6.5297 | -102.4109
6.5334 | -102.4327
6.537 | -102.458
6.5407 | -102.4868
6.5444 | -102.5109
6.548 | -102.5232
6.5517 | -102.5348
6.5553 | -102.553
6.559 | -102.5725
6.5627 | -102.5899
6.5663 | -102.6029
6.57 | -102.6146
6.5736 | -102.629
6.5773 | -102.6378
6.581 | -102.639
6.5846 | -102.6404
6.5883 | -102.6403
6.5919 | -102.6383
6.5956 | -102.6357
6.5993 | -102.6291
6.6029 | -102.6285
6.6066 | -102.6325
6.6102 | -102.637
6.6139 | -102.6488
6.6176 | -102.6638
6.6212 | -102.6777
6.6249 | -102.6849
6.6285 | -102.6959
6.6322 | -102.6987
6.6359 | -102.6934
6.6395 | -102.6913
6.6432 | -102.6817
6.6468 | -102.67
6.6505 | -102.6583
6.6542 | -102.6414
6.6578 | -102.6261
6.6615 | -102.6072
6.6651 | -102.5889
6.6688 | -102.5774
6.6725 | -102.5764
6.6761 | -102.5851
6.6798 | -102.6042
6.6834 | -102.6307
6.6871 | -102.6698
6.6908 | -102.716
6.6944 | -102.7599
6.6981 | -102.8079
6.7017 | -102.849
6.7054 | -102.8793
6.7091 | -102.8999
6.7127 | -102.9089
6.7164 | -102.9022
6.72 | -102.8869
6.7237 | -102.863
6.7274 | -102.8282
6.731 | -102.7943
6.7347 | -102.7612
6.7383 | -102.7213
6.742 | -102.6821
6.7457 | -102.6496
6.7493 | -102.6167
6.753 | -102.5954
6.7566 | -102.5826
6.7603 | -102.5782
6.764 | -102.5766
6.7676 | -102.5773
6.7713 | -102.5778
6.7749 | -102.5864
6.7786 | -102.6025
6.7823 | -102.6179
6.7859 | -102.6314
6.7896 | -102.6419
6.7932 | -102.653
6.7969 | -102.6541
6.8006 | -102.6517
6.8042 | -102.6502
6.8079 | -102.6416
6.8115 | -102.6299
6.8152 | -102.6266
6.8189 | -102.6291
6.8225 | -102.6373
6.8262 | -102.6467
6.8298 | -102.6682
6.8335 | -102.6984
6.8372 | -102.7373
6.8408 | -102.7862
6.8445 | -102.8395
6.8481 | -102.8966
6.8518 | -102.9537
6.8555 | -103.0068
6.8591 | -103.0562
6.8628 | -103.0752
6.8665 | -103.1129
6.8701 | -103.1413
6.8738 | -103.1614
6.8774 | -103.1765
6.8811 | -103.1876
6.8848 | -103.1994
6.8884 | -103.2104
6.8921 | -103.2181
6.8957 | -103.2263
6.8994 | -103.231
6.9031 | -103.2235
6.9067 | -103.2123
6.9104 | -103.1974
6.914 | -103.1754
6.9177 | -103.1467
6.9214 | -103.1101
6.925 | -103.072
6.9287 | -103.0366
6.9323 | -102.9985
6.936 | -102.9631
6.9397 | -102.9262
6.9433 | -102.8927
6.947 | -102.8604
6.9506 | -102.8369
6.9543 | -102.8205
6.958 | -102.8151
6.9616 | -102.8126
6.9653 | -102.8206
6.9689 | -102.837
6.9726 | -102.852
6.9763 | -102.8749
6.9799 | -102.8997
6.9836 | -102.9217
6.9872 | -102.9393
6.9909 | -102.9565
6.9946 | -102.97
6.9982 | -102.9867
7.0019 | -102.9921
7.0055 | -102.9976
7.0092 | -103.0018
7.0129 | -103.0005
7.0165 | -102.9973
7.0202 | -102.9925
7.0238 | -102.9803
7.0275 | -102.9689
7.0312 | -102.9546
7.0348 | -102.9313
7.0385 | -102.9079
7.0421 | -102.887
7.0458 | -102.8646
7.0495 | -102.8362
7.0531 | -102.8109
7.0568 | -102.7883
7.0604 | -102.7676
7.0641 | -102.7492
7.0678 | -102.7373
7.0714 | -102.7192
7.0751 | -102.7001
7.0787 | -102.6785
7.0824 | -102.6531
7.0861 | -102.6252
7.0897 | -102.5962
7.0934 | -102.5686
7.097 | -102.5361
7.1007 | -102.5072
7.1044 | -102.4763
7.108 | -102.4511
7.1117 | -102.4238
7.1153 | -102.3997
7.119 | -102.376
7.1227 | -102.3527
7.1263 | -102.3327
7.13 | -102.3258
7.1336 | -102.3258
7.1373 | -102.3358
7.141 | -102.3608
7.1446 | -102.3977
7.1483 | -102.4453
7.1519 | -102.508
7.1556 | -102.58
7.1593 | -102.6589
7.1629 | -102.7425
7.1666 | -102.8305
7.1702 | -102.9232
7.1739 | -103.0138
7.1776 | -103.0995
7.1812 | -103.1732
7.1849 | -103.236
7.1885 | -103.2829
7.1922 | -103.3168
7.1959 | -103.3341
7.1995 | -103.3345
7.2032 | -103.3203
7.2068 | -103.291
7.2105 | -103.2547
7.2142 | -103.2097
7.2178 | -103.1563
7.2215 | -103.0935
7.2251 | -103.0316
7.2288 | -102.9711
7.2325 | -102.9129
7.2361 | -102.8534
7.2398 | -102.8009
7.2434 | -102.7561
7.2471 | -102.7248
7.2508 | -102.7054
7.2544 | -102.6966
7.2581 | -102.7049
7.2617 | -102.7243
7.2654 | -102.7465
7.2691 | -102.7731
7.2727 | -102.8044
7.2764 | -102.8364
7.28 | -102.8727
7.2837 | -102.9109
7.2874 | -102.9518
7.291 | -102.9988
7.2947 | -103.0423
7.2983 | -103.0861
7.302 | -103.1349
7.3057 | -103.1873
7.3093 | -103.243
7.313 | -103.2949
7.3166 | -103.3463
7.3203 | -103.4013
7.324 | -103.4555
7.3276 | -103.5085
7.3313 | -103.5527
7.335 | -103.5869
7.3386 | -103.6105
7.3423 | -103.6267
7.3459 | -103.6308
7.3496 | -103.6302
7.3533 | -103.6196
7.3569 | -103.6019
7.3606 | -103.5822
7.3642 | -103.5564
7.3679 | -103.5348
7.3716 | -103.5182
7.3752 | -103.5056
7.3789 | -103.4975
7.3825 | -103.495
7.3862 | -103.4994
7.3899 | -103.509
7.3935 | -103.5104
7.3972 | -103.5068
7.4008 | -103.4953
7.4045 | -103.4875
7.4082 | -103.4739
7.4118 | -103.4486
7.4155 | -103.4151
7.4191 | -103.3752
7.4228 | -103.3298
7.4265 | -103.2811
7.4301 | -103.2294
7.4338 | -103.178
7.4374 | -103.1253
7.4411 | -103.0794
7.4448 | -103.0412
7.4484 | -103.0102
7.4521 | -102.9909
7.4557 | -102.9865
7.4594 | -102.9929
7.4631 | -103.0131
7.4667 | -103.0467
7.4704 | -103.0931
7.474 | -103.1476
7.4777 | -103.2102
7.4814 | -103.2772
7.485 | -103.3402
7.4887 | -103.4023
7.4923 | -103.4608
7.496 | -103.5052
7.4997 | -103.5367
7.5033 | -103.5577
7.507 | -103.5606
7.5106 | -103.5532
7.5143 | -103.532
7.518 | -103.4956
7.5216 | -103.4513
7.5253 | -103.4012
7.5289 | -103.3452
7.5326 | -103.2892
7.5363 | -103.2344
7.5399 | -103.1821
7.5436 | -103.1319
7.5472 | -103.0877
7.5509 | -103.0505
7.5546 | -103.0201
7.5582 | -102.9939
7.5619 | -102.9693
7.5655 | -102.9446
7.5692 | -102.9211
7.5729 | -102.9029
7.5765 | -102.8909
7.5802 | -102.8861
7.5838 | -102.8865
7.5875 | -102.8987
7.5912 | -102.9186
7.5948 | -102.9537
7.5985 | -103.0038
7.6021 | -103.0626
7.6058 | -103.13
7.6095 | -103.2116
7.6131 | -103.3008
7.6168 | -103.3944
7.6204 | -103.4857
7.6241 | -103.5725
7.6278 | -103.6537
7.6314 | -103.7271
7.6351 | -103.7887
7.6387 | -103.8399
7.6424 | -103.8845
7.6461 | -103.9085
7.6497 | -103.9183
7.6534 | -103.9159
7.657 | -103.9063
7.6607 | -103.8871
7.6644 | -103.8526
7.668 | -103.8084
7.6717 | -103.7623
7.6753 | -103.7185
7.679 | -103.6733
7.6827 | -103.6296
7.6863 | -103.5863
7.69 | -103.5485
7.6936 | -103.5184
7.6973 | -103.4926
7.701 | -103.4675
7.7046 | -103.4433
7.7083 | -103.4207
7.7119 | -103.4024
7.7156 | -103.3814
7.7193 | -103.3579
7.7229 | -103.3323
7.7266 | -103.3047
7.7302 | -103.2735
7.7339 | -103.247
7.7376 | -103.22
7.7412 | -103.1899
7.7449 | -103.1605
7.7485 | -103.133
7.7522 | -103.1087
7.7559 | -103.0815
7.7595 | -103.0618
7.7632 | -103.0435
7.7668 | -103.0309
7.7705 | -103.0258
7.7742 | -103.026
7.7778 | -103.0349
7.7815 | -103.0459
7.7851 | -103.0618
7.7888 | -103.0821
7.7925 | -103.1112
7.7961 | -103.1438
7.7998 | -103.1789
7.8034 | -103.2155
7.8071 | -103.254
7.8108 | -103.298
7.8144 | -103.3488
7.8181 | -103.4024
7.8217 | -103.4517
7.8254 | -103.5025
7.8291 | -103.562
7.8327 | -103.6232
7.8364 | -103.6807
7.8401 | -103.7413
7.8437 | -103.7906
7.8474 | -103.8279
7.851 | -103.8607
7.8547 | -103.889
7.8584 | -103.905
7.862 | -103.9134
7.8657 | -103.9188
7.8693 | -103.9215
7.873 | -103.9155
7.8767 | -103.9016
7.8803 | -103.8843
7.884 | -103.8694
7.8876 | -103.8623
7.8913 | -103.8591
7.895 | -103.8581
7.8986 | -103.8594
7.9023 | -103.8629
7.9059 | -103.8707
7.9096 | -103.886
7.9133 | -103.9107
7.9169 | -103.9441
7.9206 | -103.9852
7.9242 | -104.0198
7.9279 | -104.0556
7.9316 | -104.0894
7.9352 | -104.1179
7.9389 | -104.1461
7.9425 | -104.1615
7.9462 | -104.166
7.9499 | -104.1678
7.9535 | -104.1573
7.9572 | -104.1321
7.9608 | -104.1008
7.9645 | -104.0567
7.9682 | -104.0066
7.9718 | -103.955
7.9755 | -103.9017
7.9791 | -103.8471
7.9828 | -103.7869
7.9865 | -103.7285
7.9901 | -103.6752
7.9938 | -103.6345
7.9974 | -103.5976
8.0011 | -103.5639
8.0048 | -103.5328
8.0084 | -103.5122
8.0121 | -103.5041
8.0157 | -103.5035
8.0194 | -103.5053
8.0231 | -103.5118
8.0267 | -103.5172
8.0304 | -103.5235
8.034 | -103.5279
8.0377 | -103.5289
8.0414 | -103.5275
8.045 | -103.5174
8.0487 | -103.5018
8.0523 | -103.4821
8.056 | -103.4626
8.0597 | -103.4424
8.0633 | -103.4201
8.067 | -103.3946
8.0706 | -103.3694
8.0743 | -103.348
8.078 | -103.3247
8.0816 | -103.3031
8.0853 | -103.2867
8.0889 | -103.2696
8.0926 | -103.2573
8.0963 | -103.2503
8.0999 | -103.2451
8.1036 | -103.2534
8.1072 | -103.2696
8.1109 | -103.2944
8.1146 | -103.3235
8.1182 | -103.3534
8.1219 | -103.3855
8.1255 | -103.4254
8.1292 | -103.4607
8.1329 | -103.4901
8.1365 | -103.5184
8.1402 | -103.5501
8.1438 | -103.5748
8.1475 | -103.5954
8.1512 | -103.6124
8.1548 | -103.6285
8.1585 | -103.6446
8.1621 | -103.662
8.1658 | -103.6785
8.1695 | -103.6966
8.1731 | -103.7118
8.1768 | -103.7286
8.1804 | -103.7435
8.1841 | -103.7613
8.1878 | -103.7837
8.1914 | -103.8062
8.1951 | -103.8288
8.1987 | -103.8536
8.2024 | -103.8775
8.2061 | -103.8957
8.2097 | -103.9107
8.2134 | -103.9249
8.217 | -103.9398
8.2207 | -103.9563
8.2244 | -103.9623
8.228 | -103.9665
8.2317 | -103.9665
8.2353 | -103.961
8.239 | -103.9598
8.2427 | -103.9573
8.2463 | -103.9444
8.25 | -103.9317
8.2536 | -103.9156
8.2573 | -103.89
8.261 | -103.8601
8.2646 | -103.8357
8.2683 | -103.8151
8.2719 | -103.7911
8.2756 | -103.7676
8.2793 | -103.7455
8.2829 | -103.7262
8.2866 | -103.713
8.2903 | -103.7025
8.2939 | -103.7001
8.2976 | -103.7022
8.3012 | -103.708
8.3049 | -103.7173
8.3086 | -103.7283
8.3122 | -103.7387
8.3159 | -103.7533
8.3195 | -103.7657
8.3232 | -103.7772
8.3269 | -103.7898
8.3305 | -103.799
8.3342 | -103.7987
8.3378 | -103.799
8.3415 | -103.801
8.3452 | -103.8035
8.3488 | -103.8058
8.3525 | -103.8057
8.3561 | -103.8062
8.3598 | -103.8049
8.3635 | -103.8046
8.3671 | -103.8083
8.3708 | -103.8078
8.3744 | -103.808
8.3781 | -103.8128
8.3818 | -103.817
8.3854 | -103.8164
8.3891 | -103.8183
8.3927 | -103.8281
8.3964 | -103.8446
8.4001 | -103.867
8.4037 | -103.8867
8.4074 | -103.9073
8.411 | -103.9373
8.4147 | -103.9672
8.4184 | -103.9959
8.422 | -104.0288
8.4257 | -104.0615
8.4293 | -104.0906
8.433 | -104.1241
8.4367 | -104.1579
8.4403 | -104.1851
8.444 | -104.1994
8.4476 | -104.2016
8.4513 | -104.2046
8.455 | -104.2032
8.4586 | -104.2052
8.4623 | -104.1971
8.4659 | -104.1814
8.4696 | -104.1615
8.4733 | -104.1374
8.4769 | -104.1096
8.4806 | -104.0754
8.4842 | -104.039
8.4879 | -104.0029
8.4916 | -103.9666
8.4952 | -103.9329
8.4989 | -103.8981
8.5025 | -103.8641
8.5062 | -103.8333
8.5099 | -103.8068
8.5135 | -103.7813
8.5172 | -103.7565
8.5208 | -103.7313
8.5245 | -103.7075
8.5282 | -103.6917
8.5318 | -103.6799
8.5355 | -103.6653
8.5391 | -103.6551
8.5428 | -103.6433
8.5465 | -103.6303
8.5501 | -103.6236
8.5538 | -103.6186
8.5574 | -103.6123
8.5611 | -103.6085
8.5648 | -103.6131
8.5684 | -103.6225
8.5721 | -103.629
8.5757 | -103.6422
8.5794 | -103.6636
8.5831 | -103.6884
8.5867 | -103.7112
8.5904 | -103.7419
8.594 | -103.7777
8.5977 | -103.8141
8.6014 | -103.8555
8.605 | -103.902
8.6087 | -103.948
8.6123 | -103.9932
8.616 | -104.0356
8.6197 | -104.0755
8.6233 | -104.1145
8.627 | -104.1493
8.6306 | -104.1734
8.6343 | -104.1949
8.638 | -104.2157
8.6416 | -104.2301
8.6453 | -104.2416
8.6489 | -104.2495
8.6526 | -104.2545
8.6563 | -104.2502
8.6599 | -104.2362
8.6636 | -104.2186
8.6672 | -104.1984
8.6709 | -104.1821
8.6746 | -104.1587
8.6782 | -104.1235
8.6819 | -104.0845
8.6855 | -104.0473
8.6892 | -104.0036
8.6929 | -103.9624
8.6965 | -103.9276
8.7002 | -103.8948
8.7038 | -103.8625
8.7075 | -103.833
8.7112 | -103.809
8.7148 | -103.7946
8.7185 | -103.7824
8.7221 | -103.7747
8.7258 | -103.7801
8.7295 | -103.7847
8.7331 | -103.783
8.7368 | -103.788
8.7404 | -103.7922
8.7441 | -103.7905
8.7478 | -103.7856
8.7514 | -103.7763
8.7551 | -103.7631
8.7587 | -103.7478
8.7624 | -103.7266
8.7661 | -103.6979
8.7697 | -103.6654
8.7734 | -103.6244
8.7771 | -103.5854
8.7807 | -103.5483
8.7844 | -103.5129
8.788 | -103.4845
8.7917 | -103.4614
8.7954 | -103.4432
8.799 | -103.4299
8.8027 | -103.4295
8.8063 | -103.4386
8.81 | -103.4572
8.8137 | -103.4853
8.8173 | -103.521
8.821 | -103.559
8.8246 | -103.604
8.8283 | -103.6537
8.832 | -103.7031
8.8356 | -103.7567
8.8393 | -103.8146
8.8429 | -103.8695
8.8466 | -103.9177
8.8503 | -103.9645
8.8539 | -104.0026
8.8576 | -104.0388
8.8612 | -104.0779
8.8649 | -104.1028
8.8686 | -104.1196
8.8722 | -104.1328
8.8759 | -104.1392
8.8795 | -104.1383
8.8832 | -104.1366
8.8869 | -104.1287
8.8905 | -104.1155
8.8942 | -104.1013
8.8978 | -104.0881
8.9015 | -104.0777
8.9052 | -104.063
8.9088 | -104.0435
8.9125 | -104.0269
8.9161 | -104.0084
8.9198 | -103.9864
8.9235 | -103.9629
8.9271 | -103.9415
8.9308 | -103.9229
8.9344 | -103.9037
8.9381 | -103.8854
8.9418 | -103.8669
8.9454 | -103.8555
8.9491 | -103.8451
8.9527 | -103.8353
8.9564 | -103.8293
8.9601 | -103.8318
8.9637 | -103.8378
8.9674 | -103.8421
8.971 | -103.8507
8.9747 | -103.8782
8.9784 | -103.8971
8.982 | -103.9193
8.9857 | -103.9426
8.9893 | -103.97
8.993 | -103.9992
8.9967 | -104.0293
9.0003 | -104.0604
9.004 | -104.0944
9.0076 | -104.1264
9.0113 | -104.1584
9.015 | -104.1875
9.0186 | -104.2156
9.0223 | -104.244
9.0259 | -104.2695
9.0296 | -104.2898
9.0333 | -104.3095
9.0369 | -104.3248
9.0406 | -104.3409
9.0442 | -104.3538
9.0479 | -104.3652
9.0516 | -104.3763
9.0552 | -104.393
9.0589 | -104.4061
9.0625 | -104.4138
9.0662 | -104.4237
9.0699 | -104.435
9.0735 | -104.4439
9.0772 | -104.4515
9.0808 | -104.4673
9.0845 | -104.4835
9.0882 | -104.5028
9.0918 | -104.5183
9.0955 | -104.529
9.0991 | -104.5423
9.1028 | -104.5603
9.1065 | -104.5759
9.1101 | -104.5914
9.1138 | -104.6017
9.1174 | -104.6075
9.1211 | -104.6145
9.1248 | -104.6189
9.1284 | -104.6208
9.1321 | -104.6266
9.1357 | -104.6315
9.1394 | -104.6315
9.1431 | -104.6324
9.1467 | -104.6331
9.1504 | -104.6342
9.154 | -104.6375
9.1577 | -104.6379
9.1614 | -104.6456
9.165 | -104.6542
9.1687 | -104.655
9.1723 | -104.659
9.176 | -104.6599
9.1797 | -104.6581
9.1833 | -104.6554
9.187 | -104.6485
9.1906 | -104.6311
9.1943 | -104.6177
9.198 | -104.6081
9.2016 | -104.5973
9.2053 | -104.5813
9.2089 | -104.5692
9.2126 | -104.5603
9.2163 | -104.5449
9.2199 | -104.53
9.2236 | -104.5185
9.2272 | -104.5077
9.2309 | -104.4953
9.2346 | -104.485
9.2382 | -104.4763
9.2419 | -104.4672
9.2456 | -104.4628
9.2492 | -104.4616
9.2529 | -104.4578
9.2565 | -104.4539
9.2602 | -104.4449
9.2639 | -104.4368
9.2675 | -104.4295
9.2712 | -104.4221
9.2748 | -104.415
9.2785 | -104.407
9.2822 | -104.3994
9.2858 | -104.3896
9.2895 | -104.3765
9.2931 | -104.3617
9.2968 | -104.3455
9.3005 | -104.3348
9.3041 | -104.3264
9.3078 | -104.3154
9.3114 | -104.3031
9.3151 | -104.2956
9.3188 | -104.2897
9.3224 | -104.2853
9.3261 | -104.2858
9.3297 | -104.2828
9.3334 | -104.2798
9.3371 | -104.2818
9.3407 | -104.2829
9.3444 | -104.2816
9.348 | -104.2771
9.3517 | -104.2752
9.3554 | -104.2727
9.359 | -104.2678
9.3627 | -104.2653
9.3663 | -104.2622
9.37 | -104.254
9.3737 | -104.244
9.3773 | -104.2335
9.381 | -104.2206
9.3846 | -104.2085
9.3883 | -104.2012
9.392 | -104.1954
9.3956 | -104.1916
9.3993 | -104.1856
9.4029 | -104.1726
9.4066 | -104.1646
9.4103 | -104.1577
9.4139 | -104.1531
9.4176 | -104.1533
9.4212 | -104.1564
9.4249 | -104.1669
9.4286 | -104.1823
9.4322 | -104.1992
9.4359 | -104.2147
9.4395 | -104.2309
9.4432 | -104.2529
9.4469 | -104.2812
9.4505 | -104.3039
9.4542 | -104.3226
9.4578 | -104.3463
9.4615 | -104.3695
9.4652 | -104.3932
9.4688 | -104.4108
9.4725 | -104.4282
9.4761 | -104.4515
9.4798 | -104.4754
9.4835 | -104.502
9.4871 | -104.5272
9.4908 | -104.5521
9.4944 | -104.5763
9.4981 | -104.6038
9.5018 | -104.6283
9.5054 | -104.6532
9.5091 | -104.6784
9.5127 | -104.6989
9.5164 | -104.7166
9.5201 | -104.7325
9.5237 | -104.7477
9.5274 | -104.7544
9.531 | -104.7552
9.5347 | -104.7605
9.5384 | -104.7637
9.542 | -104.7701
9.5457 | -104.7676
9.5493 | -104.7676
9.553 | -104.7641
9.5567 | -104.7543
9.5603 | -104.7474
9.564 | -104.7388
9.5676 | -104.7327
9.5713 | -104.7252
9.575 | -104.7218
9.5786 | -104.7221
9.5823 | -104.7218
9.5859 | -104.7237
9.5896 | -104.7302
9.5933 | -104.7427
9.5969 | -104.7453
9.6006 | -104.7457
9.6042 | -104.7473
9.6079 | -104.7552
9.6116 | -104.7595
9.6152 | -104.7593
9.6189 | -104.7595
9.6225 | -104.7542
9.6262 | -104.7462
9.6299 | -104.7365
9.6335 | -104.7264
9.6372 | -104.7108
9.6408 | -104.7008
9.6445 | -104.689
9.6482 | -104.6767
9.6518 | -104.6649
9.6555 | -104.6571
9.6591 | -104.6525
9.6628 | -104.6458
9.6665 | -104.6395
9.6701 | -104.6366
9.6738 | -104.6356
9.6774 | -104.6335
9.6811 | -104.6342
9.6848 | -104.6397
9.6884 | -104.6514
9.6921 | -104.6612
9.6958 | -104.6788
9.6994 | -104.7033
9.7031 | -104.7284
9.7067 | -104.7542
9.7104 | -104.7784
9.7141 | -104.8083
9.7177 | -104.8378
9.7214 | -104.8677
9.725 | -104.8961
9.7287 | -104.9216
9.7324 | -104.9448
9.736 | -104.9625
9.7397 | -104.9814
9.7433 | -104.9918
9.747 | -104.9938
9.7507 | -104.9899
9.7543 | -104.984
9.758 | -104.9716
9.7616 | -104.9487
9.7653 | -104.9279
9.769 | -104.9001
9.7726 | -104.8662
9.7763 | -104.8307
9.7799 | -104.7937
9.7836 | -104.7514
9.7873 | -104.7092
9.7909 | -104.6674
9.7946 | -104.6236
9.7982 | -104.5803
9.8019 | -104.5411
9.8056 | -104.5062
9.8092 | -104.4735
9.8129 | -104.4469
9.8165 | -104.4257
9.8202 | -104.4067
9.8239 | -104.3932
9.8275 | -104.379
9.8312 | -104.3754
9.8348 | -104.384
9.8385 | -104.3935
9.8422 | -104.407
9.8458 | -104.4267
9.8495 | -104.4524
9.8531 | -104.4835
9.8568 | -104.5185
9.8605 | -104.5519
9.8641 | -104.5825
9.8678 | -104.6137
9.8714 | -104.6412
9.8751 | -104.6637
9.8788 | -104.6838
9.8824 | -104.709
9.8861 | -104.7329
9.8897 | -104.7534
9.8934 | -104.7687
9.8971 | -104.7754
9.9007 | -104.7777
9.9044 | -104.7799
9.908 | -104.7743
9.9117 | -104.7666
9.9154 | -104.7528
9.919 | -104.7298
9.9227 | -104.7025
9.9263 | -104.6705
9.93 | -104.6393
9.9337 | -104.6079
9.9373 | -104.5779
9.941 | -104.5521
9.9446 | -104.5283
9.9483 | -104.5087
9.952 | -104.4974
9.9556 | -104.4913
9.9593 | -104.4897
9.9629 | -104.4952
9.9666 | -104.5071
9.9703 | -104.5258
9.9739 | -104.5424
9.9776 | -104.5593
9.9812 | -104.5838
9.9849 | -104.6101
9.9886 | -104.6361
9.9922 | -104.6681
9.9959 | -104.6981
9.9995 | -104.7261
10.0032 | -104.761
10.0069 | -104.7978
10.0105 | -104.8295
10.0142 | -104.8591
10.0178 | -104.8896
10.0215 | -104.9169
10.0252 | -104.9339
10.0288 | -104.9476
10.0325 | -104.9465
10.0361 | -104.9544
10.0398 | -104.9606
10.0435 | -104.9672
10.0471 | -104.9724
10.0508 | -104.9716
10.0544 | -104.9678
10.0581 | -104.9623
10.0618 | -104.9558
10.0654 | -104.949
10.0691 | -104.9397
10.0727 | -104.9305
10.0764 | -104.9234
10.0801 | -104.9179
10.0837 | -104.9132
10.0874 | -104.9057
10.091 | -104.8957
10.0947 | -104.8896
10.0984 | -104.8851
10.102 | -104.8828
10.1057 | -104.8834
10.1093 | -104.8842
10.113 | -104.8834
10.1167 | -104.8839
10.1203 | -104.8843
10.124 | -104.8846
10.1276 | -104.8818
10.1313 | -104.8782
10.135 | -104.8769
10.1386 | -104.8707
10.1423 | -104.8679
10.1459 | -104.8628
10.1496 | -104.8535
10.1533 | -104.8425
10.1569 | -104.8347
10.1606 | -104.831
10.1642 | -104.8268
10.1679 | -104.8225
10.1716 | -104.8198
10.1752 | -104.8176
10.1789 | -104.8227
10.1826 | -104.8309
10.1862 | -104.8368
10.1899 | -104.8451
10.1935 | -104.8545
10.1972 | -104.8653
10.2009 | -104.8796
10.2045 | -104.895
10.2082 | -104.9115
10.2118 | -104.9298
10.2155 | -104.9497
10.2192 | -104.9715
10.2228 | -104.9891
10.2265 | -104.9991
10.2301 | -105.002
10.2338 | -105.0062
10.2375 | -105.0094
10.2411 | -105.009
10.2448 | -105.0019
10.2484 | -104.9894
10.2521 | -104.9717
10.2558 | -104.9485
10.2594 | -104.9259
10.2631 | -104.9057
10.2667 | -104.8836
10.2704 | -104.8581
10.2741 | -104.8327
10.2777 | -104.8072
10.2814 | -104.7823
10.285 | -104.7584
10.2887 | -104.7337
10.2924 | -104.7049
10.296 | -104.6868
10.2997 | -104.6751
10.3033 | -104.6659
10.307 | -104.6585
10.3107 | -104.6515
10.3143 | -104.6459
10.318 | -104.6452
10.3216 | -104.6496
10.3253 | -104.6563
10.329 | -104.6639
10.3326 | -104.6758
10.3363 | -104.6918
10.3399 | -104.7079
10.3436 | -104.7271
10.3473 | -104.7464
10.3509 | -104.7655
10.3546 | -104.7846
10.3582 | -104.8028
10.3619 | -104.819
10.3656 | -104.8346
10.3692 | -104.8451
10.3729 | -104.8563
10.3765 | -104.8677
10.3802 | -104.8776
10.3839 | -104.8801
10.3875 | -104.8848
10.3912 | -104.889
10.3948 | -104.8849
10.3985 | -104.8764
10.4022 | -104.8673
10.4058 | -104.8583
10.4095 | -104.8426
10.4131 | -104.8228
10.4168 | -104.8041
10.4205 | -104.7863
10.4241 | -104.7673
10.4278 | -104.7509
10.4314 | -104.7374
10.4351 | -104.7237
10.4388 | -104.7123
10.4424 | -104.7072
10.4461 | -104.7016
10.4497 | -104.6984
10.4534 | -104.7
10.4571 | -104.705
10.4607 | -104.7143
10.4644 | -104.7286
10.468 | -104.7469
10.4717 | -104.7674
10.4754 | -104.7923
10.479 | -104.8175
10.4827 | -104.845
10.4863 | -104.8784
10.49 | -104.9142
10.4937 | -104.9521
10.4973 | -104.9932
10.501 | -105.0331
10.5046 | -105.0745
10.5083 | -105.1171
10.512 | -105.1581
10.5156 | -105.1922
10.5193 | -105.2309
10.5229 | -105.2706
10.5266 | -105.3082
10.5303 | -105.3463
10.5339 | -105.3768
10.5376 | -105.3976
10.5412 | -105.4156
10.5449 | -105.4385
10.5486 | -105.4594
10.5522 | -105.4779
10.5559 | -105.4919
10.5595 | -105.4945
10.5632 | -105.5027
10.5669 | -105.5073
10.5705 | -105.5086
10.5742 | -105.5044
10.5778 | -105.4955
10.5815 | -105.4812
10.5852 | -105.4637
10.5888 | -105.4473
10.5925 | -105.4298
10.5961 | -105.4057
10.5998 | -105.3773
10.6035 | -105.3454
10.6071 | -105.3095
10.6108 | -105.2729
10.6144 | -105.2368
10.6181 | -105.1993
10.6218 | -105.1597
10.6254 | -105.1222
10.6291 | -105.0878
10.6327 | -105.05
10.6364 | -105.016
10.6401 | -104.9848
10.6437 | -104.9546
10.6474 | -104.928
10.651 | -104.9033
10.6547 | -104.8814
10.6584 | -104.865
10.662 | -104.8531
10.6657 | -104.8466
10.6694 | -104.8458
10.673 | -104.8474
10.6767 | -104.8507
10.6803 | -104.8563
10.684 | -104.8655
10.6877 | -104.8774
10.6913 | -104.8961
10.695 | -104.919
10.6986 | -104.938
10.7023 | -104.9575
10.706 | -104.9757
10.7096 | -104.9951
10.7133 | -105.0122
10.7169 | -105.0277
10.7206 | -105.0398
10.7243 | -105.0483
10.7279 | -105.0543
10.7316 | -105.061
10.7352 | -105.0648
10.7389 | -105.0647
10.7426 | -105.0623
10.7462 | -105.0549
10.7499 | -105.0465
10.7535 | -105.0374
10.7572 | -105.0266
10.7609 | -105.0199
10.7645 | -105.0157
10.7682 | -105.0109
10.7718 | -105.0067
10.7755 | -105.0051
10.7792 | -105.0026
10.7828 | -104.999
10.7865 | -104.9978
10.7901 | -105.0006
10.7938 | -105.0019
10.7975 | -105.0009
10.8011 | -105.005
10.8048 | -105.0098
10.8084 | -105.0154
10.8121 | -105.0249
10.8158 | -105.0336
10.8194 | -105.0446
10.8231 | -105.0551
10.8267 | -105.0661
10.8304 | -105.0806
10.8341 | -105.0964
10.8377 | -105.1137
10.8414 | -105.1311
10.845 | -105.1533
10.8487 | -105.1682
10.8524 | -105.18
10.856 | -105.1902
10.8597 | -105.1989
10.8633 | -105.2072
10.867 | -105.2116
10.8707 | -105.2169
10.8743 | -105.2213
10.878 | -105.2246
10.8816 | -105.2225
10.8853 | -105.2141
10.889 | -105.2076
10.8926 | -105.2004
10.8963 | -105.1866
10.8999 | -105.1715
10.9036 | -105.1586
10.9073 | -105.144
10.9109 | -105.1301
10.9146 | -105.1142
10.9182 | -105.0957
10.9219 | -105.0754
10.9256 | -105.0525
10.9292 | -105.0316
10.9329 | -105.011
10.9365 | -104.9919
10.9402 | -104.9723
10.9439 | -104.9512
10.9475 | -104.9325
10.9512 | -104.9129
10.9548 | -104.8945
10.9585 | -104.879
10.9622 | -104.862
10.9658 | -104.8479
10.9695 | -104.8344
10.9731 | -104.8234
10.9768 | -104.8151
10.9805 | -104.8114
10.9841 | -104.8084
10.9878 | -104.8079
10.9914 | -104.8121
10.9951 | -104.8207
10.9988 | -104.8338
11.0024 | -104.8472
11.0061 | -104.8693
11.0097 | -104.8971
11.0134 | -104.9309
11.0171 | -104.9654
11.0207 | -105.0005
11.0244 | -105.0425
11.028 | -105.0854
11.0317 | -105.1284
11.0354 | -105.1699
11.039 | -105.2144
11.0427 | -105.2595
11.0463 | -105.2952
11.05 | -105.3292
11.0537 | -105.3632
11.0573 | -105.3891
11.061 | -105.4092
11.0646 | -105.4246
11.0683 | -105.4346
11.072 | -105.4393
11.0756 | -105.4339
11.0793 | -105.4259
11.0829 | -105.4151
11.0866 | -105.3884
11.0903 | -105.3721
11.0939 | -105.3535
11.0976 | -105.3345
11.1012 | -105.3169
11.1049 | -105.2984
11.1086 | -105.2773
11.1122 | -105.2589
11.1159 | -105.2432
11.1196 | -105.2267
11.1232 | -105.2131
11.1269 | -105.2061
11.1305 | -105.1998
11.1342 | -105.1974
11.1379 | -105.198
11.1415 | -105.2025
11.1452 | -105.2164
11.1488 | -105.23
11.1525 | -105.2445
11.1562 | -105.2655
11.1598 | -105.2901
11.1635 | -105.3142
11.1671 | -105.3391
11.1708 | -105.3657
11.1745 | -105.394
11.1781 | -105.4264
11.1818 | -105.4581
11.1854 | -105.4853
11.1891 | -105.5105
11.1928 | -105.5372
11.1964 | -105.5634
11.2001 | -105.5849
11.2037 | -105.6026
11.2074 | -105.6203
11.2111 | -105.6325
11.2147 | -105.6439
11.2184 | -105.6524
11.222 | -105.6579
11.2257 | -105.6642
11.2294 | -105.6738
11.233 | -105.681
11.2367 | -105.6907
11.2403 | -105.6992
11.244 | -105.7042
11.2477 | -105.7092
11.2513 | -105.715
11.255 | -105.7215
11.2586 | -105.7262
11.2623 | -105.7283
11.266 | -105.7284
11.2696 | -105.7319
11.2733 | -105.7355
11.2769 | -105.7343
11.2806 | -105.7356
11.2843 | -105.738
11.2879 | -105.7376
11.2916 | -105.7328
11.2952 | -105.7316
11.2989 | -105.7266
11.3026 | -105.7196
11.3062 | -105.7116
11.3099 | -105.7011
11.3135 | -105.6926
11.3172 | -105.6866
11.3209 | -105.6737
11.3245 | -105.659
11.3282 | -105.6497
11.3318 | -105.6371
11.3355 | -105.6226
11.3392 | -105.6079
11.3428 | -105.592
11.3465 | -105.5776
11.3501 | -105.5652
11.3538 | -105.5538
11.3575 | -105.5429
11.3611 | -105.5333
11.3648 | -105.5243
11.3684 | -105.5205
11.3721 | -105.5174
11.3758 | -105.5144
11.3794 | -105.5145
11.3831 | -105.5176
11.3867 | -105.5232
11.3904 | -105.5295
11.3941 | -105.5315
11.3977 | -105.5362
11.4014 | -105.5458
11.405 | -105.5584
11.4087 | -105.5727
11.4124 | -105.5867
11.416 | -105.5986
11.4197 | -105.6112
11.4233 | -105.6272
11.427 | -105.6429
11.4307 | -105.6565
11.4343 | -105.663
11.438 | -105.6709
11.4416 | -105.6777
11.4453 | -105.6755
11.449 | -105.6705
11.4526 | -105.6653
11.4563 | -105.6573
11.4599 | -105.6431
11.4636 | -105.6302
11.4673 | -105.6155
11.4709 | -105.5981
11.4746 | -105.5714
11.4782 | -105.5432
11.4819 | -105.5129
11.4856 | -105.484
11.4892 | -105.4529
11.4929 | -105.4206
11.4965 | -105.3912
11.5002 | -105.36
11.5039 | -105.3282
11.5075 | -105.3003
11.5112 | -105.2725
11.5148 | -105.2482
11.5185 | -105.226
11.5222 | -105.2027
11.5258 | -105.182
11.5295 | -105.1665
11.5331 | -105.1503
11.5368 | -105.1388
11.5405 | -105.1319
11.5441 | -105.125
11.5478 | -105.1218
11.5514 | -105.118
11.5551 | -105.1173
11.5588 | -105.1208
11.5624 | -105.1266
11.5661 | -105.1318
11.5697 | -105.1372
11.5734 | -105.1435
11.5771 | -105.1518
11.5807 | -105.1627
11.5844 | -105.1737
11.5881 | -105.1815
11.5917 | -105.1923
11.5954 | -105.2058
11.599 | -105.2191
11.6027 | -105.2291
11.6064 | -105.2398
11.61 | -105.2533
11.6137 | -105.2671
11.6173 | -105.2806
11.621 | -105.2965
11.6247 | -105.3116
11.6283 | -105.325
11.632 | -105.3367
11.6356 | -105.3472
11.6393 | -105.355
11.643 | -105.3692
11.6466 | -105.3816
11.6503 | -105.3949
11.6539 | -105.4066
11.6576 | -105.4192
11.6613 | -105.4314
11.6649 | -105.4369
11.6686 | -105.4423
11.6722 | -105.4448
11.6759 | -105.4431
11.6796 | -105.4432
11.6832 | -105.4411
11.6869 | -105.4366
11.6905 | -105.4322
11.6942 | -105.4248
11.6979 | -105.4143
11.7015 | -105.4048
11.7052 | -105.3924
11.7088 | -105.3784
11.7125 | -105.3632
11.7162 | -105.3489
11.7198 | -105.3344
11.7235 | -105.3186
11.7271 | -105.3033
11.7308 | -105.2905
11.7345 | -105.2779
11.7381 | -105.2648
11.7418 | -105.2533
11.7454 | -105.2395
11.7491 | -105.226
11.7528 | -105.2155
11.7564 | -105.207
11.7601 | -105.1945
11.7637 | -105.1838
11.7674 | -105.1762
11.7711 | -105.17
11.7747 | -105.1644
11.7784 | -105.1593
11.782 | -105.1541
11.7857 | -105.1509
11.7894 | -105.1454
11.793 | -105.1415
11.7967 | -105.1383
11.8003 | -105.1384
11.804 | -105.1395
11.8077 | -105.1423
11.8113 | -105.1473
11.815 | -105.1537
11.8186 | -105.1626
11.8223 | -105.1735
11.826 | -105.1904
11.8296 | -105.2086
11.8333 | -105.226
11.8369 | -105.2508
11.8406 | -105.2788
11.8443 | -105.3117
11.8479 | -105.3473
11.8516 | -105.3788
11.8552 | -105.4093
11.8589 | -105.442
11.8626 | -105.4726
11.8662 | -105.5013
11.8699 | -105.5288
11.8735 | -105.5551
11.8772 | -105.5839
11.8809 | -105.6133
11.8845 | -105.6434
11.8882 | -105.6687
11.8918 | -105.6903
11.8955 | -105.7108
11.8992 | -105.7305
11.9028 | -105.7461
11.9065 | -105.7591
11.9101 | -105.771
11.9138 | -105.7819
11.9175 | -105.7912
11.9211 | -105.7997
11.9248 | -105.8043
11.9284 | -105.8026
11.9321 | -105.8008
11.9358 | -105.7947
11.9394 | -105.787
11.9431 | -105.7789
11.9467 | -105.7684
11.9504 | -105.7547
11.9541 | -105.7376
11.9577 | -105.7227
11.9614 | -105.7093
11.965 | -105.6905
11.9687 | -105.6707
11.9724 | -105.6493
11.976 | -105.6305
11.9797 | -105.6115
11.9833 | -105.593
11.987 | -105.5733
11.9907 | -105.5514
11.9943 | -105.5289
11.998 | -105.5067
12.0016 | -105.487
12.0053 | -105.4707
12.009 | -105.45
12.0126 | -105.4292
12.0163 | -105.4058
12.0199 | -105.3841
12.0236 | -105.3666
12.0273 | -105.3495
12.0309 | -105.3334
12.0346 | -105.3183
12.0382 | -105.3051
12.0419 | -105.2898
12.0456 | -105.2732
12.0492 | -105.2556
12.0529 | -105.2407
12.0565 | -105.2301
12.0602 | -105.2224
12.0639 | -105.2183
12.0675 | -105.2141
12.0712 | -105.2097
12.0748 | -105.2087
12.0785 | -105.2121
12.0822 | -105.2173
12.0858 | -105.22
12.0895 | -105.2213
12.0932 | -105.2219
12.0968 | -105.2249
12.1005 | -105.236
12.1041 | -105.2474
12.1078 | -105.2554
12.1115 | -105.2618
12.1151 | -105.2683
12.1188 | -105.2759
12.1224 | -105.2848
12.1261 | -105.2949
12.1298 | -105.3019
12.1334 | -105.3082
12.1371 | -105.3154
12.1407 | -105.3202
12.1444 | -105.3248
12.1481 | -105.3303
12.1517 | -105.3339
12.1554 | -105.3385
12.159 | -105.3432
12.1627 | -105.3483
12.1664 | -105.3575
12.17 | -105.37
12.1737 | -105.3828
12.1773 | -105.3963
12.181 | -105.4092
12.1847 | -105.4288
12.1883 | -105.4526
12.192 | -105.4788
12.1956 | -105.5026
12.1993 | -105.5249
12.203 | -105.5468
12.2066 | -105.5693
12.2103 | -105.5897
12.2139 | -105.6073
12.2176 | -105.6259
12.2213 | -105.6431
12.2249 | -105.657
12.2286 | -105.6678
12.2322 | -105.6746
12.2359 | -105.6777
12.2396 | -105.6785
12.2432 | -105.6783
12.2469 | -105.6747
12.2505 | -105.6647
12.2542 | -105.6512
12.2579 | -105.6352
12.2615 | -105.6191
12.2652 | -105.5997
12.2688 | -105.5794
12.2725 | -105.5547
12.2762 | -105.5253
12.2798 | -105.4963
12.2835 | -105.4714
12.2871 | -105.4463
12.2908 | -105.4207
12.2945 | -105.3992
12.2981 | -105.3797
12.3018 | -105.3605
12.3054 | -105.3438
12.3091 | -105.3322
12.3128 | -105.3259
12.3164 | -105.325
12.3201 | -105.3233
12.3237 | -105.3249
12.3274 | -105.3317
12.3311 | -105.3393
12.3347 | -105.3514
12.3384 | -105.3676
12.342 | -105.3858
12.3457 | -105.4023
12.3494 | -105.4191
12.353 | -105.4368
12.3567 | -105.4563
12.3603 | -105.4787
12.364 | -105.4972
12.3677 | -105.513
12.3713 | -105.5291
12.375 | -105.5485
12.3786 | -105.5685
12.3823 | -105.5876
12.386 | -105.6032
12.3896 | -105.6194
12.3933 | -105.6364
12.3969 | -105.6505
12.4006 | -105.6653
12.4043 | -105.68
12.4079 | -105.6927
12.4116 | -105.7024
12.4152 | -105.7139
12.4189 | -105.7235
12.4226 | -105.7339
12.4262 | -105.7436
12.4299 | -105.7558
12.4335 | -105.7672
12.4372 | -105.7757
12.4409 | -105.7803
12.4445 | -105.7912
12.4482 | -105.7996
12.4518 | -105.8092
12.4555 | -105.8185
12.4592 | -105.8282
12.4628 | -105.8359
12.4665 | -105.8422
12.4701 | -105.8493
12.4738 | -105.8515
12.4775 | -105.8515
12.4811 | -105.8527
12.4848 | -105.8544
12.4884 | -105.8545
12.4921 | -105.8506
12.4958 | -105.8475
12.4994 | -105.8464
12.5031 | -105.8436
12.5067 | -105.8427
12.5104 | -105.8393
12.5141 | -105.8359
12.5177 | -105.8331
12.5214 | -105.8314
12.525 | -105.8341
12.5287 | -105.8356
12.5324 | -105.8403
12.536 | -105.8463
12.5397 | -105.8553
12.5434 | -105.8689
12.547 | -105.8832
12.5507 | -105.8964
12.5543 | -105.9084
12.558 | -105.9236
12.5617 | -105.9406
12.5653 | -105.9593
12.569 | -105.9808
12.5726 | -105.9983
12.5763 | -106.015
12.58 | -106.03
12.5836 | -106.0449
12.5873 | -106.0598
12.5909 | -106.0702
12.5946 | -106.0799
12.5983 | -106.0914
12.6019 | -106.0939
12.6056 | -106.0972
12.6092 | -106.096
12.6129 | -106.0919
12.6166 | -106.093
12.6202 | -106.095
12.6239 | -106.0855
12.6275 | -106.0726
12.6312 | -106.0596
12.6349 | -106.0471
12.6385 | -106.0357
12.6422 | -106.0239
12.6458 | -106.0126
12.6495 | -105.9992
12.6532 | -105.9883
12.6568 | -105.979
12.6605 | -105.9689
12.6641 | -105.9564
12.6678 | -105.9446
12.6715 | -105.9405
12.6751 | -105.931
12.6788 | -105.9233
12.6824 | -105.9169
12.6861 | -105.9143
12.6898 | -105.9111
12.6934 | -105.9114
12.6971 | -105.9124
12.7007 | -105.9123
12.7044 | -105.913
12.7081 | -105.916
12.7117 | -105.9199
12.7154 | -105.923
12.719 | -105.9271
12.7227 | -105.9298
12.7264 | -105.9349
12.73 | -105.9385
12.7337 | -105.9444
12.7373 | -105.9507
12.741 | -105.9557
12.7447 | -105.9577
12.7483 | -105.9607
12.752 | -105.9616
12.7556 | -105.9604
12.7593 | -105.9599
12.763 | -105.9578
12.7666 | -105.9565
12.7703 | -105.9529
12.7739 | -105.9537
12.7776 | -105.9493
12.7813 | -105.9423
12.7849 | -105.9348
12.7886 | -105.927
12.7922 | -105.9177
12.7959 | -105.9053
12.7996 | -105.8897
12.8032 | -105.874
12.8069 | -105.8616
12.8105 | -105.8499
12.8142 | -105.8377
12.8179 | -105.8223
12.8215 | -105.8087
12.8252 | -105.8014
12.8288 | -105.7972
12.8325 | -105.7948
12.8362 | -105.7929
12.8398 | -105.7909
12.8435 | -105.7904
12.8471 | -105.7936
12.8508 | -105.7982
12.8545 | -105.8079
12.8581 | -105.8197
12.8618 | -105.8372
12.8654 | -105.8554
12.8691 | -105.8746
12.8728 | -105.8933
12.8764 | -105.9159
12.8801 | -105.9412
12.8837 | -105.9674
12.8874 | -105.9975
12.8911 | -106.0274
12.8947 | -106.0533
12.8984 | -106.0809
12.902 | -106.1071
12.9057 | -106.1319
12.9094 | -106.1548
12.913 | -106.1751
12.9167 | -106.2015
12.9203 | -106.2263
12.924 | -106.2453
12.9277 | -106.2631
12.9313 | -106.2845
12.935 | -106.3043
12.9386 | -106.3222
12.9423 | -106.3446
12.946 | -106.3648
12.9496 | -106.3833
12.9533 | -106.3998
12.9569 | -106.4178
12.9606 | -106.4328
12.9643 | -106.45
12.9679 | -106.4642
12.9716 | -106.4803
12.9752 | -106.4931
12.9789 | -106.508
12.9826 | -106.5226
12.9862 | -106.5338
12.9899 | -106.5432
12.9935 | -106.5498
12.9972 | -106.5527
13.0009 | -106.549
13.0045 | -106.5464
13.0082 | -106.5396
13.0119 | -106.5271
13.0155 | -106.5128
13.0192 | -106.4976
13.0228 | -106.4831
13.0265 | -106.4702
13.0302 | -106.4525
13.0338 | -106.43
13.0375 | -106.4039
13.0411 | -106.3738
13.0448 | -106.3415
13.0485 | -106.3066
13.0521 | -106.2724
13.0558 | -106.2408
13.0594 | -106.2089
13.0631 | -106.1751
13.0668 | -106.1395
13.0704 | -106.1049
13.0741 | -106.0712
13.0777 | -106.0403
13.0814 | -106.0121
13.0851 | -105.9834
13.0887 | -105.9546
13.0924 | -105.9299
13.096 | -105.9081
13.0997 | -105.8883
13.1034 | -105.8682
13.107 | -105.8497
13.1107 | -105.8346
13.1143 | -105.8224
13.118 | -105.8094
13.1217 | -105.7995
13.1253 | -105.7935
13.129 | -105.7878
13.1326 | -105.784
13.1363 | -105.784
13.14 | -105.7858
13.1436 | -105.7922
13.1473 | -105.8014
13.1509 | -105.8093
13.1546 | -105.8196
13.1583 | -105.8308
13.1619 | -105.8434
13.1656 | -105.8541
13.1692 | -105.8659
13.1729 | -105.8787
13.1766 | -105.8923
13.1802 | -105.9049
13.1839 | -105.9186
13.1875 | -105.9343
13.1912 | -105.9492
13.1949 | -105.9603
13.1985 | -105.959
13.2022 | -105.9649
13.2058 | -105.9699
13.2095 | -105.9749
13.2132 | -105.9785
13.2168 | -105.9784
13.2205 | -105.9767
13.2241 | -105.9736
13.2278 | -105.9746
13.2315 | -105.973
13.2351 | -105.9683
13.2388 | -105.9638
13.2424 | -105.958
13.2461 | -105.9489
13.2498 | -105.941
13.2534 | -105.9332
13.2571 | -105.927
13.2607 | -105.9237
13.2644 | -105.9198
13.2681 | -105.9168
13.2717 | -105.9145
13.2754 | -105.9106
13.279 | -105.9064
13.2827 | -105.903
13.2864 | -105.9026
13.29 | -105.9021
13.2937 | -105.8993
13.2973 | -105.9004
13.301 | -105.9046
13.3047 | -105.9055
13.3083 | -105.9073
13.312 | -105.908
13.3156 | -105.9098
13.3193 | -105.9116
13.323 | -105.9124
13.3266 | -105.9094
13.3303 | -105.9049
13.3339 | -105.8955
13.3376 | -105.8816
13.3413 | -105.871
13.3449 | -105.8589
13.3486 | -105.8451
13.3522 | -105.8278
13.3559 | -105.8109
13.3596 | -105.7913
13.3632 | -105.7713
13.3669 | -105.7519
13.3705 | -105.7322
13.3742 | -105.7132
13.3779 | -105.6934
13.3815 | -105.6721
13.3852 | -105.6547
13.3888 | -105.6399
13.3925 | -105.6273
13.3962 | -105.6173
13.3998 | -105.6071
13.4035 | -105.5971
13.4071 | -105.5906
13.4108 | -105.5877
13.4145 | -105.5895
13.4181 | -105.5938
13.4218 | -105.5976
13.4254 | -105.6048
13.4291 | -105.6136
13.4328 | -105.6233
13.4364 | -105.6363
13.4401 | -105.6531
13.4437 | -105.6695
13.4474 | -105.6858
13.4511 | -105.7042
13.4547 | -105.7196
13.4584 | -105.7366
13.462 | -105.7543
13.4657 | -105.7717
13.4694 | -105.7917
13.473 | -105.8126
13.4767 | -105.8271
13.4803 | -105.8411
13.484 | -105.8581
13.4877 | -105.8773
13.4913 | -105.8949
13.495 | -105.908
13.4986 | -105.9222
13.5023 | -105.9382
13.506 | -105.9508
13.5096 | -105.9626
13.5133 | -105.9772
13.517 | -105.9935
13.5206 | -106.0106
13.5243 | -106.0232
13.5279 | -106.0391
13.5316 | -106.0565
13.5353 | -106.0667
13.5389 | -106.0759
13.5426 | -106.0871
13.5462 | -106.0982
13.5499 | -106.1044
13.5536 | -106.1095
13.5572 | -106.1126
13.5609 | -106.1109
13.5645 | -106.11
13.5682 | -106.1094
13.5719 | -106.1084
13.5755 | -106.1069
13.5792 | -106.1008
13.5828 | -106.0934
13.5865 | -106.0868
13.5902 | -106.0764
13.5938 | -106.0686
13.5975 | -106.0627
13.6011 | -106.0548
13.6048 | -106.0485
13.6085 | -106.0409
13.6121 | -106.0343
13.6158 | -106.0288
13.6194 | -106.0251
13.6231 | -106.0241
13.6268 | -106.0223
13.6304 | -106.0182
13.6341 | -106.0159
13.6377 | -106.0182
13.6414 | -106.022
13.6451 | -106.0269
13.6487 | -106.0331
13.6524 | -106.0364
13.656 | -106.0417
13.6597 | -106.0497
13.6634 | -106.058
13.667 | -106.0659
13.6707 | -106.076
13.6743 | -106.0831
13.678 | -106.0882
13.6817 | -106.096
13.6853 | -106.1002
13.689 | -106.1058
13.6926 | -106.1118
13.6963 | -106.1165
13.7 | -106.1232
13.7036 | -106.1258
13.7073 | -106.1253
13.7109 | -106.1199
13.7146 | -106.1149
13.7183 | -106.1129
13.7219 | -106.1074
13.7256 | -106.0896
13.7292 | -106.0821
13.7329 | -106.0768
13.7366 | -106.0669
13.7402 | -106.0528
13.7439 | -106.0397
13.7475 | -106.0295
13.7512 | -106.0192
13.7549 | -106.0075
13.7585 | -105.9995
13.7622 | -105.992
13.7658 | -105.9843
13.7695 | -105.9803
13.7732 | -105.9799
13.7768 | -105.9806
13.7805 | -105.9828
13.7841 | -105.9877
13.7878 | -105.995
13.7915 | -106.0016
13.7951 | -106.0107
13.7988 | -106.0213
13.8024 | -106.0321
13.8061 | -106.0451
13.8098 | -106.0598
13.8134 | -106.0756
13.8171 | -106.0908
13.8207 | -106.1051
13.8244 | -106.1209
13.8281 | -106.1376
13.8317 | -106.1512
13.8354 | -106.1631
13.839 | -106.1764
13.8427 | -106.1847
13.8464 | -106.1929
13.85 | -106.2032
13.8537 | -106.2104
13.8573 | -106.2174
13.861 | -106.223
13.8647 | -106.2247
13.8683 | -106.2259
13.872 | -106.2268
13.8756 | -106.2294
13.8793 | -106.2284
13.883 | -106.227
13.8866 | -106.2263
13.8903 | -106.2213
13.8939 | -106.2156
13.8976 | -106.2112
13.9013 | -106.2092
13.9049 | -106.2105
13.9086 | -106.2143
13.9122 | -106.216
13.9159 | -106.218
13.9196 | -106.2216
13.9232 | -106.2287
13.9269 | -106.2383
13.9305 | -106.2411
13.9342 | -106.2442
13.9379 | -106.2504
13.9415 | -106.2599
13.9452 | -106.269
13.9488 | -106.2763
13.9525 | -106.2834
13.9562 | -106.2913
13.9598 | -106.2918
13.9635 | -106.287
13.9672 | -106.2857
13.9708 | -106.2841
13.9745 | -106.2819
13.9781 | -106.2807
13.9818 | -106.2764
13.9855 | -106.2674
13.9891 | -106.2573
13.9928 | -106.2504
13.9964 | -106.2401
14.0001 | -106.2294
14.0038 | -106.2191
14.0074 | -106.2078
14.0111 | -106.1898
14.0147 | -106.1662
14.0184 | -106.1473
14.0221 | -106.1314
14.0257 | -106.1148
14.0294 | -106.0966
14.033 | -106.0786
14.0367 | -106.0624
14.0404 | -106.0476
14.044 | -106.0341
14.0477 | -106.0224
14.0513 | -106.012
14.055 | -106.004
14.0587 | -105.9973
14.0623 | -105.9925
14.066 | -105.9916
14.0696 | -105.9924
14.0733 | -105.9949
14.077 | -105.9937
14.0806 | -105.9946
14.0843 | -106.0029
14.0879 | -106.0109
14.0916 | -106.0196
14.0953 | -106.0291
14.0989 | -106.0363
14.1026 | -106.0458
14.1062 | -106.0534
14.1099 | -106.0602
14.1136 | -106.0697
14.1172 | -106.0781
14.1209 | -106.0865
14.1245 | -106.0954
14.1282 | -106.1044
14.1319 | -106.1158
14.1355 | -106.128
14.1392 | -106.1382
14.1428 | -106.1457
14.1465 | -106.1561
14.1502 | -106.1688
14.1538 | -106.1819
14.1575 | -106.1934
14.1611 | -106.2111
14.1648 | -106.2317
14.1685 | -106.2434
14.1721 | -106.2563
14.1758 | -106.2694
14.1794 | -106.2821
14.1831 | -106.2953
14.1868 | -106.3106
14.1904 | -106.3254
14.1941 | -106.3349
14.1977 | -106.3437
14.2014 | -106.3546
14.2051 | -106.3642
14.2087 | -106.3717
14.2124 | -106.3818
14.216 | -106.3935
14.2197 | -106.3976
14.2234 | -106.3991
14.227 | -106.3972
14.2307 | -106.3957
14.2343 | -106.3944
14.238 | -106.3905
14.2417 | -106.3849
14.2453 | -106.379
14.249 | -106.3718
14.2526 | -106.3665
14.2563 | -106.3582
14.26 | -106.3505
14.2636 | -106.3417
14.2673 | -106.3353
14.2709 | -106.3304
14.2746 | -106.3251
14.2783 | -106.322
14.2819 | -106.3218
14.2856 | -106.3221
14.2892 | -106.3177
14.2929 | -106.3174
14.2966 | -106.3192
14.3002 | -106.3197
14.3039 | -106.3231
14.3075 | -106.3273
14.3112 | -106.3343
14.3149 | -106.3449
14.3185 | -106.3556
14.3222 | -106.3696
14.3258 | -106.3842
14.3295 | -106.401
14.3332 | -106.4194
14.3368 | -106.4384
14.3405 | -106.4566
14.3441 | -106.478
14.3478 | -106.4979
14.3515 | -106.5163
14.3551 | -106.5355
14.3588 | -106.557
14.3624 | -106.5769
14.3661 | -106.5951
14.3698 | -106.6103
14.3734 | -106.6293
14.3771 | -106.6447
14.3807 | -106.656
14.3844 | -106.6667
14.3881 | -106.6803
14.3917 | -106.691
14.3954 | -106.6971
14.399 | -106.701
14.4027 | -106.704
14.4064 | -106.7057
14.41 | -106.7065
14.4137 | -106.707
14.4173 | -106.7072
14.421 | -106.7051
14.4247 | -106.7032
14.4283 | -106.7003
14.432 | -106.6996
14.4357 | -106.6991
14.4393 | -106.6942
14.443 | -106.686
14.4466 | -106.6812
14.4503 | -106.6802
14.454 | -106.68
14.4576 | -106.6748
14.4613 | -106.6714
14.4649 | -106.6705
14.4686 | -106.6709
14.4723 | -106.6711
14.4759 | -106.6707
14.4796 | -106.6702
14.4832 | -106.6704
14.4869 | -106.6712
14.4906 | -106.6734
14.4942 | -106.6734
14.4979 | -106.6757
14.5015 | -106.6774
14.5052 | -106.6786
14.5089 | -106.6818
14.5125 | -106.684
14.5162 | -106.6828
14.5198 | -106.6814
14.5235 | -106.6827
14.5272 | -106.6826
14.5308 | -106.6826
14.5345 | -106.6849
14.5381 | -106.6874
14.5418 | -106.6925
14.5455 | -106.7002
14.5491 | -106.7095
14.5528 | -106.7179
14.5564 | -106.7228
14.5601 | -106.7273
14.5638 | -106.7315
14.5674 | -106.7339
14.5711 | -106.7409
14.5747 | -106.7462
14.5784 | -106.7519
14.5821 | -106.7554
14.5857 | -106.7543
14.5894 | -106.753
14.593 | -106.7542
14.5967 | -106.7504
14.6004 | -106.7472
14.604 | -106.7452
14.6077 | -106.7425
14.6113 | -106.7357
14.615 | -106.723
14.6187 | -106.7106
14.6223 | -106.6975
14.626 | -106.6836
14.6296 | -106.6674
14.6333 | -106.6492
14.637 | -106.6276
14.6406 | -106.6063
14.6443 | -106.5833
14.6479 | -106.5634
14.6516 | -106.5466
14.6553 | -106.5245
14.6589 | -106.4982
14.6626 | -106.4742
14.6662 | -106.4553
14.6699 | -106.4376
14.6736 | -106.4211
14.6772 | -106.405
14.6809 | -106.3912
14.6845 | -106.3784
14.6882 | -106.3674
14.6919 | -106.3593
14.6955 | -106.3523
14.6992 | -106.3465
14.7028 | -106.3431
14.7065 | -106.342
14.7102 | -106.34
14.7138 | -106.3401
14.7175 | -106.3369
14.7211 | -106.3367
14.7248 | -106.339
14.7285 | -106.3433
14.7321 | -106.3448
14.7358 | -106.3471
14.7394 | -106.3468
14.7431 | -106.3481
14.7468 | -106.3495
14.7504 | -106.3475
14.7541 | -106.3423
14.7577 | -106.3385
14.7614 | -106.3339
14.7651 | -106.334
14.7687 | -106.3327
14.7724 | -106.3287
14.776 | -106.324
14.7797 | -106.3177
14.7834 | -106.3137
14.787 | -106.3092
14.7907 | -106.3051
14.7943 | -106.298
14.798 | -106.2902
14.8017 | -106.2842
14.8053 | -106.2785
14.809 | -106.2747
14.8126 | -106.2718
14.8163 | -106.2714
14.82 | -106.2725
14.8236 | -106.2733
14.8273 | -106.2744
14.8309 | -106.2756
14.8346 | -106.2804
14.8383 | -106.2857
14.8419 | -106.2938
14.8456 | -106.3049
14.8492 | -106.3185
14.8529 | -106.3293
14.8566 | -106.3398
14.8602 | -106.3503
14.8639 | -106.3573
14.8675 | -106.3643
14.8712 | -106.374
14.8749 | -106.3832
14.8785 | -106.3901
14.8822 | -106.3932
14.8858 | -106.3978
14.8895 | -106.402
14.8932 | -106.4057
14.8968 | -106.4077
14.9005 | -106.4075
14.9041 | -106.4059
14.9078 | -106.4025
14.9115 | -106.3976
14.9151 | -106.3944
14.9188 | -106.3907
14.9224 | -106.3887
14.9261 | -106.385
14.9298 | -106.3824
14.9334 | -106.3779
14.9371 | -106.3726
14.9408 | -106.3651
14.9444 | -106.3575
14.9481 | -106.3533
14.9517 | -106.3506
14.9554 | -106.3472
14.9591 | -106.3436
14.9627 | -106.344
14.9664 | -106.3466
14.97 | -106.3465
14.9737 | -106.3481
14.9774 | -106.3514
14.981 | -106.3568
14.9847 | -106.3627
14.9883 | -106.3667
14.992 | -106.3696
14.9957 | -106.3726
14.9993 | -106.3809
15.003 | -106.3893
15.0066 | -106.3995
15.0103 | -106.4143
15.014 | -106.4324
15.0176 | -106.4471
15.0213 | -106.4573
15.0249 | -106.4685
15.0286 | -106.4795
15.0323 | -106.49
15.0359 | -106.5012
15.0396 | -106.5136
15.0432 | -106.5279
15.0469 | -106.5405
15.0506 | -106.5474
15.0542 | -106.5509
15.0579 | -106.5533
15.0615 | -106.5533
15.0652 | -106.5555
15.0689 | -106.5611
15.0725 | -106.5671
15.0762 | -106.5683
15.0798 | -106.5697
15.0835 | -106.573
15.0872 | -106.5754
15.0908 | -106.575
15.0945 | -106.576
15.0981 | -106.5805
15.1018 | -106.5809
15.1055 | -106.5786
15.1091 | -106.5793
15.1128 | -106.5812
15.1164 | -106.5838
15.1201 | -106.5866
15.1238 | -106.588
15.1274 | -106.591
15.1311 | -106.5921
15.1347 | -106.5899
15.1384 | -106.5868
15.1421 | -106.5857
15.1457 | -106.5863
15.1494 | -106.5852
15.153 | -106.5822
15.1567 | -106.5806
15.1604 | -106.5818
15.164 | -106.5796
15.1677 | -106.5773
15.1713 | -106.5734
15.175 | -106.5732
15.1787 | -106.5693
15.1823 | -106.5683
15.186 | -106.568
15.1896 | -106.5647
15.1933 | -106.5604
15.197 | -106.5566
15.2006 | -106.5573
15.2043 | -106.5613
15.2079 | -106.5672
15.2116 | -106.5736
15.2153 | -106.5775
15.2189 | -106.5814
15.2226 | -106.5865
15.2262 | -106.5906
15.2299 | -106.5958
15.2336 | -106.6031
15.2372 | -106.611
15.2409 | -106.6182
15.2445 | -106.6274
15.2482 | -106.6395
15.2519 | -106.6533
15.2555 | -106.6657
15.2592 | -106.6751
15.2628 | -106.6826
15.2665 | -106.6904
15.2702 | -106.7018
15.2738 | -106.7146
15.2775 | -106.7246
15.2811 | -106.7283
15.2848 | -106.73
15.2885 | -106.7275
15.2921 | -106.7281
15.2958 | -106.7287
15.2994 | -106.7263
15.3031 | -106.7196
15.3068 | -106.7112
15.3104 | -106.692
15.3141 | -106.6799
15.3177 | -106.667
15.3214 | -106.6549
15.3251 | -106.6427
15.3287 | -106.626
15.3324 | -106.6101
15.336 | -106.5922
15.3397 | -106.5711
15.3434 | -106.5524
15.347 | -106.5334
15.3507 | -106.5167
15.3543 | -106.4995
15.358 | -106.4807
15.3617 | -106.4637
15.3653 | -106.4464
15.369 | -106.4319
15.3726 | -106.4192
15.3763 | -106.4054
15.38 | -106.3922
15.3836 | -106.3815
15.3873 | -106.3733
15.391 | -106.366
15.3946 | -106.3577
15.3983 | -106.3527
15.4019 | -106.3472
15.4056 | -106.3446
15.4093 | -106.3451
15.4129 | -106.3469
15.4166 | -106.3478
15.4202 | -106.3468
15.4239 | -106.3472
15.4276 | -106.3487
15.4312 | -106.349
15.4349 | -106.3495
15.4385 | -106.3524
15.4422 | -106.3546
15.4459 | -106.3566
15.4495 | -106.3582
15.4532 | -106.3615
15.4568 | -106.3647
15.4605 | -106.3668
15.4642 | -106.3673
15.4678 | -106.3664
15.4715 | -106.3663
15.4751 | -106.3683
15.4788 | -106.3717
15.4825 | -106.377
15.4861 | -106.3816
15.4898 | -106.3835
15.4934 | -106.3818
15.4971 | -106.3777
15.5008 | -106.3763
15.5044 | -106.3767
15.5081 | -106.3789
15.5117 | -106.3811
15.5154 | -106.3823
15.5191 | -106.381
15.5227 | -106.3823
15.5264 | -106.3852
15.53 | -106.3866
15.5337 | -106.3885
15.5374 | -106.3917
15.541 | -106.3953
15.5447 | -106.3981
15.5483 | -106.404
15.552 | -106.4084
15.5557 | -106.4145
15.5593 | -106.4226
15.563 | -106.4322
15.5666 | -106.4404
15.5703 | -106.45
15.574 | -106.4574
15.5776 | -106.4657
15.5813 | -106.4772
15.5849 | -106.4897
15.5886 | -106.5005
15.5923 | -106.5111
15.5959 | -106.5249
15.5996 | -106.536
15.6032 | -106.5462
15.6069 | -106.5585
15.6106 | -106.5705
15.6142 | -106.5837
15.6179 | -106.5939
15.6215 | -106.6068
15.6252 | -106.6192
15.6289 | -106.63
15.6325 | -106.638
15.6362 | -106.6402
15.6398 | -106.646
15.6435 | -106.6528
15.6472 | -106.6569
15.6508 | -106.6594
15.6545 | -106.661
15.6581 | -106.6623
15.6618 | -106.664
15.6655 | -106.6649
15.6691 | -106.665
15.6728 | -106.6636
15.6764 | -106.6585
15.6801 | -106.6521
15.6838 | -106.646
15.6874 | -106.6357
15.6911 | -106.627
15.6947 | -106.6184
15.6984 | -106.606
15.7021 | -106.594
15.7057 | -106.582
15.7094 | -106.5668
15.713 | -106.5471
15.7167 | -106.5338
15.7204 | -106.5245
15.724 | -106.5127
15.7277 | -106.4957
15.7313 | -106.4784
15.735 | -106.46
15.7387 | -106.4419
15.7423 | -106.4248
15.746 | -106.4066
15.7496 | -106.387
15.7533 | -106.3671
15.757 | -106.3458
15.7606 | -106.3249
15.7643 | -106.3037
15.7679 | -106.2854
15.7716 | -106.2665
15.7753 | -106.2508
15.7789 | -106.2321
15.7826 | -106.215
15.7862 | -106.2017
15.7899 | -106.1879
15.7936 | -106.1738
15.7972 | -106.1643
15.8009 | -106.1571
15.8045 | -106.1513
15.8082 | -106.1468
15.8119 | -106.144
15.8155 | -106.1415
15.8192 | -106.1412
15.8228 | -106.1449
15.8265 | -106.1486
15.8302 | -106.1534
15.8338 | -106.162
15.8375 | -106.1716
15.8411 | -106.1862
15.8448 | -106.198
15.8485 | -106.2087
15.8521 | -106.2213
15.8558 | -106.2327
15.8595 | -106.248
15.8631 | -106.2678
15.8668 | -106.2831
15.8704 | -106.2968
15.8741 | -106.3106
15.8778 | -106.3248
15.8814 | -106.3415
15.8851 | -106.3561
15.8887 | -106.3668
15.8924 | -106.3759
15.8961 | -106.3823
15.8997 | -106.3861
15.9034 | -106.3928
15.907 | -106.4015
15.9107 | -106.4067
15.9144 | -106.412
15.918 | -106.414
15.9217 | -106.4134
15.9253 | -106.4139
15.929 | -106.412
15.9327 | -106.4087
15.9363 | -106.4077
15.94 | -106.4055
15.9436 | -106.3971
15.9473 | -106.3878
15.951 | -106.3774
15.9546 | -106.3706
15.9583 | -106.3644
15.9619 | -106.3566
15.9656 | -106.3497
15.9693 | -106.3446
15.9729 | -106.3395
15.9766 | -106.3353
15.9802 | -106.3328
15.9839 | -106.3307
15.9876 | -106.3276
15.9912 | -106.325
15.9949 | -106.3218
15.9985 | -106.3182
16.0022 | -106.3164
16.0059 | -106.3188
16.0095 | -106.3214
16.0132 | -106.322
16.0168 | -106.3245
16.0205 | -106.3262
16.0242 | -106.329
16.0278 | -106.3329
16.0315 | -106.338
16.0351 | -106.3416
16.0388 | -106.3484
16.0425 | -106.3564
16.0461 | -106.3616
16.0498 | -106.3695
16.0534 | -106.378
16.0571 | -106.3859
16.0608 | -106.3906
16.0644 | -106.3976
16.0681 | -106.4086
16.0717 | -106.4173
16.0754 | -106.4259
16.0791 | -106.4336
16.0827 | -106.4389
16.0864 | -106.4467
16.09 | -106.4524
16.0937 | -106.4564
16.0974 | -106.4613
16.101 | -106.4665
16.1047 | -106.4725
16.1083 | -106.4778
16.112 | -106.4824
16.1157 | -106.4868
16.1193 | -106.4932
16.123 | -106.4996
16.1266 | -106.5052
16.1303 | -106.5125
16.134 | -106.5184
16.1376 | -106.521
16.1413 | -106.5278
16.1449 | -106.5361
16.1486 | -106.5427
16.1523 | -106.5485
16.1559 | -106.556
16.1596 | -106.5655
16.1632 | -106.5746
16.1669 | -106.5842
16.1706 | -106.5983
16.1742 | -106.6107
16.1779 | -106.6218
16.1815 | -106.6346
16.1852 | -106.6443
16.1889 | -106.6524
16.1925 | -106.6629
16.1962 | -106.6733
16.1998 | -106.6839
16.2035 | -106.696
16.2072 | -106.7074
16.2108 | -106.7174
16.2145 | -106.7262
16.2181 | -106.7355
16.2218 | -106.7463
16.2255 | -106.754
16.2291 | -106.7631
16.2328 | -106.7684
16.2364 | -106.7717
16.2401 | -106.7749
16.2438 | -106.7797
16.2474 | -106.7855
16.2511 | -106.7898
16.2547 | -106.7901
16.2584 | -106.7904
16.2621 | -106.7897
16.2657 | -106.7921
16.2694 | -106.7938
16.273 | -106.7945
16.2767 | -106.7971
16.2804 | -106.7972
16.284 | -106.7982
16.2877 | -106.7971
16.2913 | -106.795
16.295 | -106.7936
16.2987 | -106.7901
16.3023 | -106.7893
16.306 | -106.7857
16.3096 | -106.7831
16.3133 | -106.7796
16.317 | -106.7765
16.3206 | -106.7756
16.3243 | -106.7748
16.328 | -106.7721
16.3316 | -106.7715
16.3353 | -106.7706
16.3389 | -106.771
16.3426 | -106.768
16.3463 | -106.7632
16.3499 | -106.7591
16.3536 | -106.7552
16.3572 | -106.7544
16.3609 | -106.7502
16.3646 | -106.7416
16.3682 | -106.7308
16.3719 | -106.7237
16.3755 | -106.7132
16.3792 | -106.7024
16.3829 | -106.6933
16.3865 | -106.6874
16.3902 | -106.6816
16.3938 | -106.6751
16.3975 | -106.6684
16.4012 | -106.6622
16.4048 | -106.654
16.4085 | -106.6494
16.4121 | -106.645
16.4158 | -106.6405
16.4195 | -106.6373
16.4231 | -106.6382
16.4268 | -106.6372
16.4304 | -106.6351
16.4341 | -106.6324
16.4378 | -106.6316
16.4414 | -106.6336
16.4451 | -106.6385
16.4487 | -106.6447
16.4524 | -106.6513
16.4561 | -106.6571
16.4597 | -106.6655
16.4634 | -106.6728
16.467 | -106.6834
16.4707 | -106.6955
16.4744 | -106.7083
16.478 | -106.719
16.4817 | -106.7303
16.4853 | -106.7438
16.489 | -106.758
16.4927 | -106.7716
16.4963 | -106.7845
16.5 | -106.7959
16.5036 | -106.8084
16.5073 | -106.8216
16.511 | -106.8323
16.5146 | -106.8429
16.5183 | -106.8543
16.5219 | -106.8682
16.5256 | -106.882
16.5293 | -106.8917
16.5329 | -106.9015
16.5366 | -106.9087
16.5402 | -106.9172
16.5439 | -106.9243
16.5476 | -106.9308
16.5512 | -106.9356
16.5549 | -106.9386
16.5585 | -106.9377
16.5622 | -106.9374
16.5659 | -106.9354
16.5695 | -106.9326
16.5732 | -106.928
16.5768 | -106.9253
16.5805 | -106.9239
16.5842 | -106.9203
16.5878 | -106.9177
16.5915 | -106.9151
16.5951 | -106.9133
16.5988 | -106.9128
16.6025 | -106.9146
16.6061 | -106.9158
16.6098 | -106.9162
16.6134 | -106.9177
16.6171 | -106.9214
16.6208 | -106.926
16.6244 | -106.9286
16.6281 | -106.9299
16.6317 | -106.9342
16.6354 | -106.9415
16.6391 | -106.9503
16.6427 | -106.9611
16.6464 | -106.9701
16.65 | -106.9802
16.6537 | -106.9927
16.6574 | -107.0061
16.661 | -107.0187
16.6647 | -107.0323
16.6683 | -107.0472
16.672 | -107.0632
16.6757 | -107.0767
16.6793 | -107.0894
16.683 | -107.1007
16.6866 | -107.1111
16.6903 | -107.1187
16.694 | -107.1256
16.6976 | -107.1325
16.7013 | -107.1391
16.7049 | -107.1405
16.7086 | -107.1394
16.7123 | -107.1388
16.7159 | -107.1389
16.7196 | -107.1374
16.7232 | -107.1326
16.7269 | -107.1242
16.7306 | -107.1165
16.7342 | -107.1075
16.7379 | -107.0978
16.7415 | -107.0861
16.7452 | -107.0727
16.7489 | -107.0613
16.7525 | -107.0504
16.7562 | -107.0384
16.7598 | -107.0205
16.7635 | -107.0055
16.7672 | -106.9903
16.7708 | -106.9763
16.7745 | -106.9631
16.7781 | -106.9497
16.7818 | -106.9333
16.7855 | -106.9196
16.7891 | -106.9076
16.7928 | -106.8953
16.7965 | -106.8854
16.8001 | -106.8772
16.8038 | -106.8684
16.8074 | -106.8578
16.8111 | -106.8492
16.8148 | -106.84
16.8184 | -106.8334
16.8221 | -106.8281
16.8257 | -106.8239
16.8294 | -106.8199
16.8331 | -106.8161
16.8367 | -106.8128
16.8404 | -106.8113
16.844 | -106.8098
16.8477 | -106.8111
16.8514 | -106.8148
16.855 | -106.8174
16.8587 | -106.8196
16.8623 | -106.8238
16.866 | -106.8297
16.8697 | -106.8349
16.8733 | -106.8412
16.877 | -106.8462
16.8806 | -106.8523
16.8843 | -106.8606
16.888 | -106.8708
16.8916 | -106.8787
16.8953 | -106.8878
16.8989 | -106.8962
16.9026 | -106.9041
16.9063 | -106.9148
16.9099 | -106.9241
16.9136 | -106.9316
16.9172 | -106.9393
16.9209 | -106.9487
16.9246 | -106.9598
16.9282 | -106.9718
16.9319 | -106.9836
16.9355 | -106.9948
16.9392 | -107.0046
16.9429 | -107.0137
16.9465 | -107.026
16.9502 | -107.0384
16.9538 | -107.0479
16.9575 | -107.0555
16.9612 | -107.062
16.9648 | -107.0707
16.9685 | -107.082
16.9721 | -107.0909
16.9758 | -107.0992
16.9795 | -107.1035
16.9831 | -107.107
16.9868 | -107.112
16.9904 | -107.1154
16.9941 | -107.1165
16.9978 | -107.1166
17.0014 | -107.1198
17.0051 | -107.1229
17.0087 | -107.1272
17.0124 | -107.1281
17.0161 | -107.1257
17.0197 | -107.1195
17.0234 | -107.1175
17.027 | -107.1146
17.0307 | -107.108
17.0344 | -107.0992
17.038 | -107.0917
17.0417 | -107.0841
17.0453 | -107.0753
17.049 | -107.0671
17.0527 | -107.0583
17.0563 | -107.0488
17.06 | -107.0362
17.0636 | -107.0225
17.0673 | -107.0079
17.071 | -106.9941
17.0746 | -106.9836
17.0783 | -106.9729
17.0819 | -106.9596
17.0856 | -106.9496
17.0893 | -106.9399
17.0929 | -106.9304
17.0966 | -106.9219
17.1002 | -106.9141
17.1039 | -106.9073
17.1076 | -106.9023
17.1112 | -106.8958
17.1149 | -106.8897
17.1185 | -106.8882
17.1222 | -106.8888
17.1259 | -106.8882
17.1295 | -106.8842
17.1332 | -106.8831
17.1368 | -106.8839
17.1405 | -106.8827
17.1442 | -106.8837
17.1478 | -106.8868
17.1515 | -106.8876
17.1551 | -106.8905
17.1588 | -106.8945
17.1625 | -106.9011
17.1661 | -106.9058
17.1698 | -106.9087
17.1734 | -106.9123
17.1771 | -106.9185
17.1808 | -106.9239
17.1844 | -106.931
17.1881 | -106.9382
17.1917 | -106.9421
17.1954 | -106.9454
17.1991 | -106.9496
17.2027 | -106.9567
17.2064 | -106.9637
17.21 | -106.9695
17.2137 | -106.9758
17.2174 | -106.9812
17.221 | -106.9874
17.2247 | -106.9987
17.2283 | -107.0091
17.232 | -107.019
17.2357 | -107.0282
17.2393 | -107.0378
17.243 | -107.0479
17.2466 | -107.0557
17.2503 | -107.0608
17.254 | -107.0663
17.2576 | -107.0742
17.2613 | -107.0842
17.2649 | -107.0961
17.2686 | -107.1045
17.2723 | -107.113
17.2759 | -107.121
17.2796 | -107.1308
17.2832 | -107.1392
17.2869 | -107.1455
17.2906 | -107.1499
17.2942 | -107.1543
17.2979 | -107.1579
17.3016 | -107.1628
17.3052 | -107.1693
17.3089 | -107.1728
17.3125 | -107.1741
17.3162 | -107.1737
17.3199 | -107.1719
17.3235 | -107.1693
17.3272 | -107.1656
17.3308 | -107.1597
17.3345 | -107.1529
17.3382 | -107.1457
17.3418 | -107.1365
17.3455 | -107.129
17.3491 | -107.1207
17.3528 | -107.1099
17.3565 | -107.0997
17.3601 | -107.0862
17.3638 | -107.0711
17.3674 | -107.0564
17.3711 | -107.0444
17.3748 | -107.036
17.3784 | -107.0263
17.3821 | -107.0162
17.3857 | -107.0073
17.3894 | -106.9976
17.3931 | -106.9858
17.3967 | -106.9762
17.4004 | -106.9648
17.404 | -106.9556
17.4077 | -106.946
17.4114 | -106.9378
17.415 | -106.9293
17.4187 | -106.9223
17.4223 | -106.9214
17.426 | -106.9163
17.4297 | -106.912
17.4333 | -106.9068
17.437 | -106.9047
17.4406 | -106.9032
17.4443 | -106.8983
17.448 | -106.8967
17.4516 | -106.8954
17.4553 | -106.8953
17.4589 | -106.896
17.4626 | -106.894
17.4663 | -106.8947
17.4699 | -106.8976
17.4736 | -106.8993
17.4772 | -106.9003
17.4809 | -106.9014
17.4846 | -106.9051
17.4882 | -106.9091
17.4919 | -106.9114
17.4955 | -106.9155
17.4992 | -106.9206
17.5029 | -106.9241
17.5065 | -106.9272
17.5102 | -106.9314
17.5138 | -106.9351
17.5175 | -106.9404
17.5212 | -106.9446
17.5248 | -106.9499
17.5285 | -106.955
17.5321 | -106.9608
17.5358 | -106.9666
17.5395 | -106.975
17.5431 | -106.9833
17.5468 | -106.9889
17.5504 | -106.9943
17.5541 | -106.9994
17.5578 | -107.0085
17.5614 | -107.0185
17.5651 | -107.0271
17.5687 | -107.0318
17.5724 | -107.039
17.5761 | -107.0488
17.5797 | -107.058
17.5834 | -107.0671
17.587 | -107.0752
17.5907 | -107.0841
17.5944 | -107.09
17.598 | -107.0984
17.6017 | -107.1078
17.6053 | -107.1145
17.609 | -107.124
17.6127 | -107.1312
17.6163 | -107.1348
17.62 | -107.1379
17.6236 | -107.1406
17.6273 | -107.1417
17.631 | -107.143
17.6346 | -107.1409
17.6383 | -107.1399
17.6419 | -107.1402
17.6456 | -107.1409
17.6493 | -107.1399
17.6529 | -107.1351
17.6566 | -107.1298
17.6602 | -107.1241
17.6639 | -107.1201
17.6676 | -107.1137
17.6712 | -107.1051
17.6749 | -107.0985
17.6785 | -107.0894
17.6822 | -107.0815
17.6859 | -107.0761
17.6895 | -107.0692
17.6932 | -107.0625
17.6968 | -107.0599
17.7005 | -107.0535
17.7042 | -107.0462
17.7078 | -107.0395
17.7115 | -107.0328
17.7151 | -107.0263
17.7188 | -107.0177
17.7225 | -107.0079
17.7261 | -107.0016
17.7298 | -106.9983
17.7334 | -106.9958
17.7371 | -106.9916
17.7408 | -106.9896
17.7444 | -106.9877
17.7481 | -106.9832
17.7518 | -106.978
17.7554 | -106.9732
17.7591 | -106.9699
17.7627 | -106.9676
17.7664 | -106.9671
17.7701 | -106.9658
17.7737 | -106.9642
17.7774 | -106.9641
17.781 | -106.9647
17.7847 | -106.9652
17.7884 | -106.9662
17.792 | -106.9669
17.7957 | -106.9698
17.7993 | -106.9741
17.803 | -106.9787
17.8067 | -106.9814
17.8103 | -106.9832
17.814 | -106.9825
17.8176 | -106.9835
17.8213 | -106.9879
17.825 | -106.9892
17.8286 | -106.989
17.8323 | -106.9906
17.8359 | -106.9922
17.8396 | -106.993
17.8433 | -106.9942
17.8469 | -106.9962
17.8506 | -106.9941
17.8542 | -106.9903
17.8579 | -106.9891
17.8616 | -106.9856
17.8652 | -106.9803
17.8689 | -106.9773
17.8725 | -106.9732
17.8762 | -106.9693
17.8799 | -106.9676
17.8835 | -106.9657
17.8872 | -106.9652
17.8908 | -106.9626
17.8945 | -106.9586
17.8982 | -106.9548
17.9018 | -106.9524
17.9055 | -106.9508
17.9091 | -106.9474
17.9128 | -106.9426
17.9165 | -106.9382
17.9201 | -106.9342
17.9238 | -106.9308
17.9274 | -106.9265
17.9311 | -106.9253
17.9348 | -106.9244
17.9384 | -106.9226
17.9421 | -106.9213
17.9457 | -106.9204
17.9494 | -106.9184
17.9531 | -106.9149
17.9567 | -106.9134
17.9604 | -106.9087
17.964 | -106.9078
17.9677 | -106.9069
17.9714 | -106.9071
17.975 | -106.9082
17.9787 | -106.9094
17.9823 | -106.9069
17.986 | -106.9033
17.9897 | -106.901
17.9933 | -106.8981
17.997 | -106.8949
18.0006 | -106.8943
18.0043 | -106.8939
18.008 | -106.8936
18.0116 | -106.8951
18.0153 | -106.8942
18.0189 | -106.8923
18.0226 | -106.8875
18.0263 | -106.883
18.0299 | -106.88
18.0336 | -106.8761
18.0372 | -106.8747
18.0409 | -106.8732
18.0446 | -106.8718
18.0482 | -106.8711
18.0519 | -106.8687
18.0555 | -106.868
18.0592 | -106.8666
18.0629 | -106.8669
18.0665 | -106.8671
18.0702 | -106.87
18.0738 | -106.8721
18.0775 | -106.8747
18.0812 | -106.8768
18.0848 | -106.8776
18.0885 | -106.8775
18.0921 | -106.8769
18.0958 | -106.8755
18.0995 | -106.8723
18.1031 | -106.8691
18.1068 | -106.8688
18.1104 | -106.8697
18.1141 | -106.8691
18.1178 | -106.8684
18.1214 | -106.8679
18.1251 | -106.8655
18.1287 | -106.863
18.1324 | -106.86
18.1361 | -106.8558
18.1397 | -106.8525
18.1434 | -106.8489
18.147 | -106.8439
18.1507 | -106.8407
18.1544 | -106.8376
18.158 | -106.8332
18.1617 | -106.8287
18.1653 | -106.8254
18.169 | -106.8222
18.1727 | -106.8165
18.1763 | -106.8122
18.18 | -106.8084
18.1836 | -106.8046
18.1873 | -106.7995
18.191 | -106.7938
18.1946 | -106.7876
18.1983 | -106.7811
18.2019 | -106.7783
18.2056 | -106.7774
18.2093 | -106.7782
18.2129 | -106.776
18.2166 | -106.7742
18.2203 | -106.7733
18.2239 | -106.7701
18.2276 | -106.7715
18.2312 | -106.7737
18.2349 | -106.7768
18.2386 | -106.7804
18.2422 | -106.7837
18.2459 | -106.7882
18.2495 | -106.7935
18.2532 | -106.8009
18.2569 | -106.808
18.2605 | -106.8132
18.2642 | -106.8225
18.2678 | -106.8308
18.2715 | -106.8356
18.2752 | -106.84
18.2788 | -106.8453
18.2825 | -106.8507
18.2861 | -106.8572
18.2898 | -106.8649
18.2935 | -106.8713
18.2971 | -106.875
18.3008 | -106.8777
18.3044 | -106.8779
18.3081 | -106.8775
18.3118 | -106.8768
18.3154 | -106.8753
18.3191 | -106.8753
18.3227 | -106.8735
18.3264 | -106.8704
18.3301 | -106.8657
18.3337 | -106.8623
18.3374 | -106.8555
18.341 | -106.8507
18.3447 | -106.8464
18.3484 | -106.8435
18.352 | -106.8424
18.3557 | -106.8407
18.3593 | -106.8401
18.363 | -106.8397
18.3667 | -106.8357
18.3703 | -106.8338
18.374 | -106.8319
18.3776 | -106.8298
18.3813 | -106.8289
18.385 | -106.8298
18.3886 | -106.8299
18.3923 | -106.8335
18.3959 | -106.8382
18.3996 | -106.8426
18.4033 | -106.8462
18.4069 | -106.8531
18.4106 | -106.8608
18.4142 | -106.8657
18.4179 | -106.8749
18.4216 | -106.8862
18.4252 | -106.8978
18.4289 | -106.9109
18.4325 | -106.9245
18.4362 | -106.9356
18.4399 | -106.9481
18.4435 | -106.9604
18.4472 | -106.9731
18.4508 | -106.9881
18.4545 | -107.0047
18.4582 | -107.0211
18.4618 | -107.0363
18.4655 | -107.051
18.4691 | -107.0632
18.4728 | -107.0762
18.4765 | -107.0891
18.4801 | -107.0922
18.4838 | -107.1043
18.4874 | -107.1157
18.4911 | -107.1215
18.4948 | -107.1272
18.4984 | -107.1317
18.5021 | -107.1362
18.5057 | -107.1426
18.5094 | -107.1491
18.5131 | -107.152
18.5167 | -107.1541
18.5204 | -107.1568
18.524 | -107.1586
18.5277 | -107.156
18.5314 | -107.1546
18.535 | -107.1535
18.5387 | -107.1511
18.5423 | -107.1471
18.546 | -107.1456
18.5497 | -107.1437
18.5533 | -107.1387
18.557 | -107.1338
18.5606 | -107.1286
18.5643 | -107.1254
18.568 | -107.1249
18.5716 | -107.121
18.5753 | -107.1167
18.5789 | -107.1124
18.5826 | -107.1053
18.5863 | -107.0998
18.5899 | -107.0946
18.5936 | -107.092
18.5972 | -107.0894
18.6009 | -107.0886
18.6046 | -107.0854
18.6082 | -107.0829
18.6119 | -107.0794
18.6155 | -107.074
18.6192 | -107.0692
18.6229 | -107.066
18.6265 | -107.0627
18.6302 | -107.0608
18.6338 | -107.0565
18.6375 | -107.0537
18.6412 | -107.0521
18.6448 | -107.051
18.6485 | -107.0485
18.6521 | -107.0465
18.6558 | -107.0418
18.6595 | -107.0377
18.6631 | -107.0354
18.6668 | -107.0332
18.6704 | -107.0299
18.6741 | -107.0268
18.6778 | -107.0251
18.6814 | -107.0225
18.6851 | -107.0189
18.6887 | -107.014
18.6924 | -107.0076
18.6961 | -107.0035
18.6997 | -107.0016
18.7034 | -106.9993
18.707 | -106.9984
18.7107 | -106.9983
18.7144 | -106.9973
18.718 | -106.9976
18.7217 | -107.0001
18.7254 | -107.001
18.729 | -107.0019
18.7327 | -107.0069
18.7363 | -107.012
18.74 | -107.0191
18.7437 | -107.028
18.7473 | -107.035
18.751 | -107.0422
18.7546 | -107.0488
18.7583 | -107.0556
18.762 | -107.0624
18.7656 | -107.0717
18.7693 | -107.0796
18.7729 | -107.0874
18.7766 | -107.0936
18.7803 | -107.1015
18.7839 | -107.1094
18.7876 | -107.1146
18.7912 | -107.121
18.7949 | -107.1273
18.7986 | -107.1334
18.8022 | -107.1419
18.8059 | -107.1494
18.8095 | -107.1555
18.8132 | -107.1589
18.8169 | -107.1591
18.8205 | -107.1594
18.8242 | -107.1609
18.8278 | -107.1628
18.8315 | -107.164
18.8352 | -107.1634
18.8388 | -107.1659
18.8425 | -107.1672
18.8461 | -107.1658
18.8498 | -107.1657
18.8535 | -107.1649
18.8571 | -107.1622
18.8608 | -107.1571
18.8644 | -107.1502
18.8681 | -107.1436
18.8718 | -107.1394
18.8754 | -107.1335
18.8791 | -107.1252
18.8827 | -107.1189
18.8864 | -107.1133
18.8901 | -107.1085
18.8937 | -107.1021
18.8974 | -107.0946
18.901 | -107.088
18.9047 | -107.0823
18.9084 | -107.0749
18.912 | -107.0663
18.9157 | -107.059
18.9193 | -107.0519
18.923 | -107.0439
18.9267 | -107.0364
18.9303 | -107.0292
18.934 | -107.0254
18.9376 | -107.0211
18.9413 | -107.0163
18.945 | -107.0112
18.9486 | -107.0049
18.9523 | -106.9992
18.9559 | -106.9917
18.9596 | -106.9858
18.9633 | -106.9805
18.9669 | -106.9755
18.9706 | -106.969
18.9742 | -106.9608
18.9779 | -106.955
18.9816 | -106.9493
18.9852 | -106.9432
18.9889 | -106.9359
18.9925 | -106.9293
18.9962 | -106.9227
18.9999 | -106.9147
19.0035 | -106.9081
19.0072 | -106.899
19.0108 | -106.8894
19.0145 | -106.8801
19.0182 | -106.8696
19.0218 | -106.8603
19.0255 | -106.8512
19.0291 | -106.8429
19.0328 | -106.8355
19.0365 | -106.8268
19.0401 | -106.819
19.0438 | -106.8105
19.0474 | -106.8038
19.0511 | -106.7974
19.0548 | -106.7914
19.0584 | -106.7862
19.0621 | -106.7813
19.0657 | -106.7778
19.0694 | -106.7751
19.0731 | -106.7723
19.0767 | -106.77
19.0804 | -106.7696
19.084 | -106.7717
19.0877 | -106.776
19.0914 | -106.7803
19.095 | -106.785
19.0987 | -106.79
19.1023 | -106.7943
19.106 | -106.7986
19.1097 | -106.8033
19.1133 | -106.8109
19.117 | -106.8199
19.1206 | -106.8301
19.1243 | -106.8383
19.128 | -106.8501
19.1316 | -106.8636
19.1353 | -106.8782
19.1389 | -106.8931
19.1426 | -106.9061
19.1463 | -106.9207
19.1499 | -106.9366
19.1536 | -106.9519
19.1572 | -106.9669
19.1609 | -106.9801
19.1646 | -106.994
19.1682 | -107.0085
19.1719 | -107.0221
19.1756 | -107.0332
19.1792 | -107.0423
19.1829 | -107.0521
19.1865 | -107.0594
19.1902 | -107.068
19.1939 | -107.0752
19.1975 | -107.079
19.2012 | -107.081
19.2048 | -107.0817
19.2085 | -107.0813
19.2122 | -107.0779
19.2158 | -107.0754
19.2195 | -107.0727
19.2231 | -107.0676
19.2268 | -107.0634
19.2305 | -107.0578
19.2341 | -107.0463
19.2378 | -107.0314
19.2414 | -107.0162
19.2451 | -107.001
19.2488 | -106.9827
19.2524 | -106.9654
19.2561 | -106.9475
19.2597 | -106.9318
19.2634 | -106.916
19.2671 | -106.8986
19.2707 | -106.8828
19.2744 | -106.867
19.278 | -106.8524
19.2817 | -106.8383
19.2854 | -106.8263
19.289 | -106.8152
19.2927 | -106.8064
19.2963 | -106.7976
19.3 | -106.7891
19.3037 | -106.7834
19.3073 | -106.7772
19.311 | -106.772
19.3146 | -106.7684
19.3183 | -106.7671
19.322 | -106.7678
19.3256 | -106.7704
19.3293 | -106.7743
19.3329 | -106.7797
19.3366 | -106.7864
19.3403 | -106.7943
19.3439 | -106.8036
19.3476 | -106.8134
19.3512 | -106.8243
19.3549 | -106.8337
19.3586 | -106.8437
19.3622 | -106.855
19.3659 | -106.8684
19.3695 | -106.8819
19.3732 | -106.8949
19.3769 | -106.9091
19.3805 | -106.9251
19.3842 | -106.939
19.3878 | -106.9501
19.3915 | -106.9618
19.3952 | -106.9734
19.3988 | -106.9862
19.4025 | -106.9949
19.4061 | -107.0008
19.4098 | -107.0032
19.4135 | -107.0038
19.4171 | -107.0057
19.4208 | -107.0074
19.4244 | -107.0083
19.4281 | -107.0063
19.4318 | -107.0037
19.4354 | -107.0015
19.4391 | -106.9995
19.4427 | -106.9966
19.4464 | -106.9915
19.4501 | -106.9856
19.4537 | -106.9799
19.4574 | -106.9719
19.461 | -106.963
19.4647 | -106.9529
19.4684 | -106.9427
19.472 | -106.9332
19.4757 | -106.9223
19.4793 | -106.9117
19.483 | -106.9008
19.4867 | -106.8908
19.4903 | -106.8793
19.494 | -106.8675
19.4976 | -106.8574
19.5013 | -106.8478
19.505 | -106.8369
19.5086 | -106.8291
19.5123 | -106.8236
19.5159 | -106.8188
19.5196 | -106.8147
19.5233 | -106.8115
19.5269 | -106.8079
19.5306 | -106.8059
19.5342 | -106.8114
19.5379 | -106.8088
19.5416 | -106.8085
19.5452 | -106.8089
19.5489 | -106.809
19.5525 | -106.8098
19.5562 | -106.8123
19.5599 | -106.817
19.5635 | -106.8219
19.5672 | -106.8279
19.5708 | -106.8338
19.5745 | -106.8397
19.5782 | -106.8451
19.5818 | -106.8477
19.5855 | -106.8539
19.5891 | -106.8609
19.5928 | -106.868
19.5965 | -106.8745
19.6001 | -106.8803
19.6038 | -106.8841
19.6074 | -106.8864
19.6111 | -106.8894
19.6148 | -106.8924
19.6184 | -106.8954
19.6221 | -106.896
19.6257 | -106.8968
19.6294 | -106.8989
19.6331 | -106.8998
19.6367 | -106.8977
19.6404 | -106.8929
19.6441 | -106.8883
19.6477 | -106.8851
19.6514 | -106.8819
19.655 | -106.8773
19.6587 | -106.8716
19.6624 | -106.8668
19.666 | -106.8633
19.6697 | -106.856
19.6733 | -106.8455
19.677 | -106.8349
19.6807 | -106.825
19.6843 | -106.8172
19.688 | -106.811
19.6916 | -106.8045
19.6953 | -106.7967
19.699 | -106.7874
19.7026 | -106.7793
19.7063 | -106.773
19.7099 | -106.7653
19.7136 | -106.7561
19.7173 | -106.7468
19.7209 | -106.7375
19.7246 | -106.7295
19.7282 | -106.7239
19.7319 | -106.7182
19.7356 | -106.7126
19.7392 | -106.708
19.7429 | -106.7025
19.7465 | -106.6984
19.7502 | -106.6937
19.7539 | -106.6888
19.7575 | -106.6826
19.7612 | -106.676
19.7648 | -106.6692
19.7685 | -106.6652
19.7722 | -106.6593
19.7758 | -106.6547
19.7795 | -106.6506
19.7831 | -106.647
19.7868 | -106.6447
19.7905 | -106.6415
19.7941 | -106.6386
19.7978 | -106.6359
19.8014 | -106.6325
19.8051 | -106.6298
19.8088 | -106.6275
19.8124 | -106.6254
19.8161 | -106.624
19.8197 | -106.6232
19.8234 | -106.6225
19.8271 | -106.6236
19.8307 | -106.6233
19.8344 | -106.6215
19.838 | -106.6189
19.8417 | -106.6161
19.8454 | -106.6139
19.849 | -106.6129
19.8527 | -106.6134
19.8563 | -106.6144
19.86 | -106.6158
19.8637 | -106.6185
19.8673 | -106.6173
19.871 | -106.6174
19.8746 | -106.6191
19.8783 | -106.6201
19.882 | -106.6201
19.8856 | -106.6222
19.8893 | -106.6264
19.8929 | -106.6308
19.8966 | -106.6357
19.9003 | -106.64
19.9039 | -106.6449
19.9076 | -106.6517
19.9112 | -106.6598
19.9149 | -106.6713
19.9186 | -106.6835
19.9222 | -106.6974
19.9259 | -106.7103
19.9295 | -106.7216
19.9332 | -106.7345
19.9369 | -106.7497
19.9405 | -106.7661
19.9442 | -106.7813
19.9478 | -106.7936
19.9515 | -106.8096
19.9552 | -106.8254
19.9588 | -106.8401
19.9625 | -106.8553
19.9661 | -106.8691
19.9698 | -106.8808
19.9735 | -106.8919
19.9771 | -106.9031
19.9808 | -106.9145
19.9844 | -106.9223
19.9881 | -106.928
19.9918 | -106.9325
19.9954 | -106.9347
19.9991 | -106.9378
20.0027 | -106.9396
20.0064 | -106.9374
20.0101 | -106.9314
20.0137 | -106.9255
20.0174 | -106.9181
20.021 | -106.9057
20.0247 | -106.8904
20.0284 | -106.8723
20.032 | -106.8549
20.0357 | -106.8364
20.0393 | -106.8158
20.043 | -106.7953
20.0467 | -106.7738
20.0503 | -106.7495
20.054 | -106.7268
20.0576 | -106.7001
20.0613 | -106.6717
20.065 | -106.6379
20.0686 | -106.6127
20.0723 | -106.5854
20.0759 | -106.5593
20.0796 | -106.534
20.0833 | -106.5063
20.0869 | -106.4797
20.0906 | -106.4537
20.0942 | -106.4275
20.0979 | -106.4012
20.1016 | -106.3767
20.1052 | -106.3551
20.1089 | -106.3337
20.1125 | -106.3125
20.1162 | -106.2961
20.1199 | -106.2813
20.1235 | -106.2691
20.1272 | -106.2573
20.1309 | -106.2437
20.1345 | -106.2313
20.1382 | -106.2211
20.1418 | -106.2127
20.1455 | -106.205
20.1492 | -106.2002
20.1528 | -106.1983
20.1565 | -106.1959
20.1601 | -106.1953
20.1638 | -106.1971
20.1675 | -106.1997
20.1711 | -106.203
20.1748 | -106.2056
20.1784 | -106.2092
20.1821 | -106.2127
20.1858 | -106.2175
20.1894 | -106.2218
20.1931 | -106.2245
20.1967 | -106.2284
20.2004 | -106.2331
20.2041 | -106.2386
20.2077 | -106.2448
20.2114 | -106.2496
20.215 | -106.2542
20.2187 | -106.2556
20.2224 | -106.2552
20.226 | -106.2555
20.2297 | -106.2563
20.2333 | -106.261
20.237 | -106.2642
20.2407 | -106.2679
20.2443 | -106.2714
20.248 | -106.2743
20.2516 | -106.2741
20.2553 | -106.2758
20.259 | -106.2775
20.2626 | -106.2785
20.2663 | -106.2783
20.2699 | -106.2806
20.2736 | -106.2824
20.2773 | -106.2851
20.2809 | -106.2889
20.2846 | -106.2893
20.2882 | -106.2928
20.2919 | -106.298
20.2956 | -106.3031
20.2992 | -106.3068
20.3029 | -106.3096
20.3065 | -106.3147
20.3102 | -106.3205
20.3139 | -106.328
20.3175 | -106.3357
20.3212 | -106.3418
20.3248 | -106.3471
20.3285 | -106.3508
20.3322 | -106.3559
20.3358 | -106.361
20.3395 | -106.3661
20.3431 | -106.3725
20.3468 | -106.3781
20.3505 | -106.3827
20.3541 | -106.3878
20.3578 | -106.3919
20.3614 | -106.397
20.3651 | -106.3996
20.3688 | -106.4002
20.3724 | -106.405
20.3761 | -106.4107
20.3797 | -106.4154
20.3834 | -106.4195
20.3871 | -106.4235
20.3907 | -106.4256
20.3944 | -106.4253
20.398 | -106.4269
20.4017 | -106.429
20.4054 | -106.4284
20.409 | -106.4284
20.4127 | -106.4262
20.4163 | -106.4229
20.42 | -106.4231
20.4237 | -106.4225
20.4273 | -106.4198
20.431 | -106.4182
20.4346 | -106.4162
20.4383 | -106.4119
20.442 | -106.4076
20.4456 | -106.4034
20.4493 | -106.3989
20.4529 | -106.3958
20.4566 | -106.3953
20.4603 | -106.3924
20.4639 | -106.3923
20.4676 | -106.3941
20.4712 | -106.3945
20.4749 | -106.394
20.4786 | -106.3963
20.4822 | -106.3998
20.4859 | -106.4016
20.4895 | -106.4026
20.4932 | -106.4046
20.4969 | -106.4068
20.5005 | -106.4111
20.5042 | -106.4161
20.5078 | -106.4241
20.5115 | -106.4319
20.5152 | -106.4395
20.5188 | -106.4479
20.5225 | -106.4537
20.5261 | -106.4603
20.5298 | -106.4677
20.5335 | -106.4757
20.5371 | -106.4824
20.5408 | -106.4913
20.5444 | -106.4992
20.5481 | -106.5074
20.5518 | -106.5132
20.5554 | -106.5208
20.5591 | -106.5271
20.5627 | -106.5323
20.5664 | -106.5394
20.5701 | -106.5464
20.5737 | -106.5544
20.5774 | -106.5588
20.5811 | -106.5617
20.5847 | -106.5659
20.5884 | -106.5717
20.592 | -106.5735
20.5957 | -106.5772
20.5994 | -106.5789
20.603 | -106.5823
20.6067 | -106.5833
20.6103 | -106.583
20.614 | -106.5828
20.6177 | -106.5823
20.6213 | -106.5822
20.625 | -106.5802
20.6286 | -106.5774
20.6323 | -106.5726
20.636 | -106.5674
20.6396 | -106.5633
20.6433 | -106.5613
20.6469 | -106.5557
20.6506 | -106.5466
20.6543 | -106.5389
20.6579 | -106.5319
20.6616 | -106.5235
20.6652 | -106.5122
20.6689 | -106.4992
20.6726 | -106.4865
20.6762 | -106.4731
20.6799 | -106.4594
20.6835 | -106.4471
20.6872 | -106.4357
20.6909 | -106.4211
20.6945 | -106.4052
20.6982 | -106.3906
20.7018 | -106.3752
20.7055 | -106.361
20.7092 | -106.348
20.7128 | -106.3332
20.7165 | -106.318
20.7201 | -106.3045
20.7238 | -106.2892
20.7275 | -106.2749
20.7311 | -106.2614
20.7348 | -106.2461
20.7384 | -106.233
20.7421 | -106.2203
20.7458 | -106.2076
20.7494 | -106.195
20.7531 | -106.184
20.7567 | -106.1719
20.7604 | -106.1589
20.7641 | -106.1494
20.7677 | -106.1437
20.7714 | -106.138
20.775 | -106.1333
20.7787 | -106.13
20.7824 | -106.1273
20.786 | -106.1242
20.7897 | -106.1222
20.7933 | -106.1235
20.797 | -106.1257
20.8007 | -106.1272
20.8043 | -106.1323
20.808 | -106.1365
20.8116 | -106.1413
20.8153 | -106.1469
20.819 | -106.1538
20.8226 | -106.1621
20.8263 | -106.1709
20.8299 | -106.1789
20.8336 | -106.1876
20.8373 | -106.1964
20.8409 | -106.2055
20.8446 | -106.2182
20.8482 | -106.2308
20.8519 | -106.2441
20.8556 | -106.2584
20.8592 | -106.2694
20.8629 | -106.279
20.8665 | -106.2899
20.8702 | -106.3011
20.8739 | -106.3126
20.8775 | -106.3209
20.8812 | -106.3289
20.8848 | -106.3359
20.8885 | -106.342
20.8922 | -106.3508
20.8958 | -106.3605
20.8995 | -106.3681
20.9031 | -106.3747
20.9068 | -106.3802
20.9105 | -106.3862
20.9141 | -106.3923
20.9178 | -106.3962
20.9214 | -106.401
20.9251 | -106.4053
20.9288 | -106.4096
20.9324 | -106.4133
20.9361 | -106.4156
20.9397 | -106.4168
20.9434 | -106.4194
20.9471 | -106.4253
20.9507 | -106.4296
20.9544 | -106.4318
20.958 | -106.4367
20.9617 | -106.4374
20.9654 | -106.4367
20.969 | -106.4361
20.9727 | -106.4362
20.9763 | -106.4378
20.98 | -106.4379
20.9837 | -106.4348
20.9873 | -106.434
20.991 | -106.4356
20.9946 | -106.4329
20.9983 | -106.4305
21.002 | -106.4311
21.0056 | -106.4309
21.0093 | -106.4311
21.0129 | -106.431
21.0166 | -106.4285
21.0203 | -106.423
21.0239 | -106.4168
21.0276 | -106.4108
21.0312 | -106.4057
21.0349 | -106.4005
21.0386 | -106.3959
21.0422 | -106.3897
21.0459 | -106.3825
21.0495 | -106.3771
21.0532 | -106.3718
21.0569 | -106.3656
21.0605 | -106.3584
21.0642 | -106.3487
21.0678 | -106.3374
21.0715 | -106.325
21.0752 | -106.3116
21.0788 | -106.301
21.0825 | -106.2908
21.0862 | -106.278
21.0898 | -106.2604
21.0935 | -106.2405
21.0971 | -106.2212
21.1008 | -106.196
21.1045 | -106.1732
21.1081 | -106.1534
21.1118 | -106.1363
21.1154 | -106.1179
21.1191 | -106.0973
21.1228 | -106.0765
21.1264 | -106.0554
21.1301 | -106.0357
21.1337 | -106.0167
21.1374 | -105.9968
21.1411 | -105.9763
21.1447 | -105.9543
21.1484 | -105.9325
21.152 | -105.9102
21.1557 | -105.8894
21.1594 | -105.871
21.163 | -105.8515
21.1667 | -105.83
21.1703 | -105.8093
21.174 | -105.7906
21.1777 | -105.7733
21.1813 | -105.7559
21.185 | -105.7399
21.1886 | -105.7255
21.1923 | -105.7095
21.196 | -105.6942
21.1996 | -105.6811
21.2033 | -105.6704
21.2069 | -105.6604
21.2106 | -105.6512
21.2143 | -105.6439
21.2179 | -105.6377
21.2216 | -105.632
21.2252 | -105.6271
21.2289 | -105.6235
21.2326 | -105.6224
21.2362 | -105.6235
21.2399 | -105.6254
21.2435 | -105.628
21.2472 | -105.6326
21.2509 | -105.6374
21.2545 | -105.6458
21.2582 | -105.6559
21.2618 | -105.6696
21.2655 | -105.6853
21.2692 | -105.7008
21.2728 | -105.7169
21.2765 | -105.7322
21.2801 | -105.7476
21.2838 | -105.7652
21.2875 | -105.7838
21.2911 | -105.8009
21.2948 | -105.82
21.2984 | -105.8391
21.3021 | -105.8574
21.3058 | -105.8752
21.3094 | -105.8916
21.3131 | -105.9097
21.3167 | -105.9291
21.3204 | -105.95
21.3241 | -105.9703
21.3277 | -105.9905
21.3314 | -106.0092
21.335 | -106.0229
21.3387 | -106.0368
21.3424 | -106.0513
21.346 | -106.067
21.3497 | -106.0807
21.3533 | -106.0928
21.357 | -106.1075
21.3607 | -106.1208
21.3643 | -106.1336
21.368 | -106.148
21.3716 | -106.1597
21.3753 | -106.1719
21.379 | -106.1807
21.3826 | -106.1894
21.3863 | -106.1978
21.3899 | -106.2065
21.3936 | -106.2108
21.3973 | -106.217
21.4009 | -106.2203
21.4046 | -106.2182
21.4082 | -106.2108
21.4119 | -106.2013
21.4156 | -106.1914
21.4192 | -106.1799
21.4229 | -106.1673
21.4265 | -106.1556
21.4302 | -106.1409
21.4339 | -106.1227
21.4375 | -106.1002
21.4412 | -106.0764
21.4448 | -106.0502
21.4485 | -106.0227
21.4522 | -105.991
21.4558 | -105.9584
21.4595 | -105.9236
21.4631 | -105.8869
21.4668 | -105.8486
21.4705 | -105.81
21.4741 | -105.7703
21.4778 | -105.7267
21.4814 | -105.6836
21.4851 | -105.637
21.4888 | -105.5882
21.4924 | -105.5399
21.4961 | -105.4918
21.4997 | -105.4435
21.5034 | -105.3937
21.5071 | -105.3441
21.5107 | -105.2954
21.5144 | -105.2479
21.518 | -105.2018
21.5217 | -105.1574
21.5254 | -105.1137
21.529 | -105.0731
21.5327 | -105.0361
21.5363 | -105.0006
21.54 | -104.9669
21.5437 | -104.9353
21.5473 | -104.9066
21.551 | -104.8805
21.5547 | -104.8599
21.5583 | -104.8424
21.562 | -104.8271
21.5656 | -104.8166
21.5693 | -104.8077
21.573 | -104.8025
21.5766 | -104.8015
21.5803 | -104.8049
21.5839 | -104.8131
21.5876 | -104.824
21.5913 | -104.8375
21.5949 | -104.8551
21.5986 | -104.8772
21.6022 | -104.9035
21.6059 | -104.9319
21.6096 | -104.9645
21.6132 | -105.0006
21.6169 | -105.0389
21.6205 | -105.0783
21.6242 | -105.1205
21.6279 | -105.1644
21.6315 | -105.2094
21.6352 | -105.2568
21.6388 | -105.308
21.6425 | -105.3584
21.6462 | -105.4104
21.6498 | -105.4631
21.6535 | -105.513
21.6571 | -105.5646
21.6608 | -105.6167
21.6645 | -105.6658
21.6681 | -105.7158
21.6718 | -105.7658
21.6754 | -105.8144
21.6791 | -105.8632
21.6828 | -105.9108
21.6864 | -105.9581
21.6901 | -106.0036
21.6937 | -106.0477
21.6974 | -106.0899
21.7011 | -106.1333
21.7047 | -106.1712
21.7084 | -106.2111
21.712 | -106.2494
21.7157 | -106.2849
21.7194 | -106.3154
21.723 | -106.3429
21.7267 | -106.3693
21.7303 | -106.3938
21.734 | -106.4148
21.7377 | -106.4356
21.7413 | -106.4567
21.745 | -106.4762
21.7486 | -106.4925
21.7523 | -106.5076
21.756 | -106.5202
21.7596 | -106.5334
21.7633 | -106.544
21.7669 | -106.5532
21.7706 | -106.5597
21.7743 | -106.5665
21.7779 | -106.5711
21.7816 | -106.5729
21.7852 | -106.5764
21.7889 | -106.5775
21.7926 | -106.5806
21.7962 | -106.5855
21.7999 | -106.5896
21.8035 | -106.5924
21.8072 | -106.5948
21.8109 | -106.5972
21.8145 | -106.5996
21.8182 | -106.6015
21.8218 | -106.6032
21.8255 | -106.6046
21.8292 | -106.6067
21.8328 | -106.6092
21.8365 | -106.6092
21.8401 | -106.6087
21.8438 | -106.6074
21.8475 | -106.6072
21.8511 | -106.6082
21.8548 | -106.6098
21.8584 | -106.6103
21.8621 | -106.6079
21.8658 | -106.6068
21.8694 | -106.6062
21.8731 | -106.6048
21.8767 | -106.6063
21.8804 | -106.609
21.8841 | -106.6101
21.8877 | -106.6092
21.8914 | -106.6084
21.895 | -106.6094
21.8987 | -106.6114
21.9024 | -106.6123
21.906 | -106.6122
21.9097 | -106.6127
21.9133 | -106.6126
21.917 | -106.6113
21.9207 | -106.6096
21.9243 | -106.6101
21.928 | -106.613
21.9316 | -106.6154
21.9353 | -106.6162
21.939 | -106.6168
21.9426 | -106.6177
21.9463 | -106.6179
21.9499 | -106.6198
21.9536 | -106.6209
21.9573 | -106.6204
21.9609 | -106.6191
21.9646 | -106.6202
21.9682 | -106.619
21.9719 | -106.6192
21.9756 | -106.6191
21.9792 | -106.6162
21.9829 | -106.6137
21.9865 | -106.6113
21.9902 | -106.6108
21.9939 | -106.6099
21.9975 | -106.6078
22.0012 | -106.6051
22.0049 | -106.6023
22.0085 | -106.6002
22.0122 | -106.5983
22.0158 | -106.5944
22.0195 | -106.5908
22.0232 | -106.5868
22.0268 | -106.5826
22.0305 | -106.5773
22.0341 | -106.5728
22.0378 | -106.5678
22.0415 | -106.5627
22.0451 | -106.5571
22.0488 | -106.5522
22.0524 | -106.5458
22.0561 | -106.5408
22.0598 | -106.5358
22.0634 | -106.5291
22.0671 | -106.5218
22.0707 | -106.5149
22.0744 | -106.5094
22.0781 | -106.5034
22.0817 | -106.4987
22.0854 | -106.4938
22.089 | -106.4897
22.0927 | -106.4862
22.0964 | -106.4821
22.1 | -106.4774
22.1037 | -106.4731
22.1073 | -106.4678
22.111 | -106.4627
22.1147 | -106.4586
22.1183 | -106.4554
22.122 | -106.4541
22.1256 | -106.4531
22.1293 | -106.4533
22.133 | -106.4513
22.1366 | -106.4463
22.1403 | -106.4445
22.1439 | -106.4445
22.1476 | -106.4457
22.1513 | -106.4473
22.1549 | -106.4506
22.1586 | -106.4542
22.1622 | -106.4552
22.1659 | -106.4575
22.1696 | -106.4604
22.1732 | -106.464
22.1769 | -106.4705
22.1805 | -106.475
22.1842 | -106.4793
22.1879 | -106.4832
22.1915 | -106.4875
22.1952 | -106.4915
22.1988 | -106.4967
22.2025 | -106.5026
22.2062 | -106.5076
22.2098 | -106.511
22.2135 | -106.5135
22.2171 | -106.517
22.2208 | -106.5214
22.2245 | -106.525
22.2281 | -106.5299
22.2318 | -106.536
22.2354 | -106.5434
22.2391 | -106.5474
22.2428 | -106.5489
22.2464 | -106.5501
22.2501 | -106.554
22.2537 | -106.5597
22.2574 | -106.5628
22.2611 | -106.5674
22.2647 | -106.5698
22.2684 | -106.5721
22.272 | -106.5748
22.2757 | -106.576
22.2794 | -106.5779
22.283 | -106.5801
22.2867 | -106.5815
22.2903 | -106.5827
22.294 | -106.5848
22.2977 | -106.5855
22.3013 | -106.5869
22.305 | -106.5863
22.3086 | -106.5848
22.3123 | -106.5853
22.316 | -106.5847
22.3196 | -106.5848
22.3233 | -106.5845
22.3269 | -106.5847
22.3306 | -106.5847
22.3343 | -106.5837
22.3379 | -106.5821
22.3416 | -106.5798
22.3452 | -106.5776
22.3489 | -106.5757
22.3526 | -106.5733
22.3562 | -106.5714
22.3599 | -106.57
22.3635 | -106.5652
22.3672 | -106.5611
22.3709 | -106.5561
22.3745 | -106.5532
22.3782 | -106.5499
22.3818 | -106.5445
22.3855 | -106.5365
22.3892 | -106.5263
22.3928 | -106.5142
22.3965 | -106.5043
22.4001 | -106.4958
22.4038 | -106.4868
22.4075 | -106.4773
22.4111 | -106.4682
22.4148 | -106.4586
22.4184 | -106.4473
22.4221 | -106.4353
22.4258 | -106.4239
22.4294 | -106.4119
22.4331 | -106.3989
22.4367 | -106.3848
22.4404 | -106.37
22.4441 | -106.3545
22.4477 | -106.3408
22.4514 | -106.328
22.455 | -106.3091
22.4587 | -106.2875
22.4624 | -106.266
22.466 | -106.246
22.4697 | -106.2243
22.4733 | -106.2025
22.477 | -106.1833
22.4807 | -106.1621
22.4843 | -106.1415
22.488 | -106.122
22.4916 | -106.1011
22.4953 | -106.0806
22.499 | -106.0612
22.5026 | -106.0434
22.5063 | -106.0251
22.51 | -106.0075
22.5136 | -105.9919
22.5173 | -105.9778
22.5209 | -105.9656
22.5246 | -105.955
22.5283 | -105.945
22.5319 | -105.9342
22.5356 | -105.9268
22.5392 | -105.9203
22.5429 | -105.9144
22.5466 | -105.9105
22.5502 | -105.9074
22.5539 | -105.9045
22.5575 | -105.9037
22.5612 | -105.9072
22.5649 | -105.9131
22.5685 | -105.9196
22.5722 | -105.9249
22.5758 | -105.9334
22.5795 | -105.9422
22.5832 | -105.9525
22.5868 | -105.9644
22.5905 | -105.9773
22.5941 | -105.9921
22.5978 | -106.0061
22.6015 | -106.0214
22.6051 | -106.0368
22.6088 | -106.0544
22.6124 | -106.0725
22.6161 | -106.0899
22.6198 | -106.1066
22.6234 | -106.1245
22.6271 | -106.1444
22.6307 | -106.1624
22.6344 | -106.1808
22.6381 | -106.1979
22.6417 | -106.2171
22.6454 | -106.2359
22.649 | -106.2524
22.6527 | -106.2701
22.6564 | -106.2885
22.66 | -106.3083
22.6637 | -106.3284
22.6673 | -106.3469
22.671 | -106.3657
22.6747 | -106.3835
22.6783 | -106.4013
22.682 | -106.4215
22.6856 | -106.4403
22.6893 | -106.4576
22.693 | -106.4732
22.6966 | -106.4886
22.7003 | -106.5055
22.7039 | -106.5173
22.7076 | -106.5354
22.7113 | -106.5521
22.7149 | -106.5697
22.7186 | -106.5823
22.7222 | -106.5925
22.7259 | -106.603
22.7296 | -106.613
22.7332 | -106.6219
22.7369 | -106.6306
22.7405 | -106.6385
22.7442 | -106.6462
22.7479 | -106.6522
22.7515 | -106.6563
22.7552 | -106.6612
22.7588 | -106.6673
22.7625 | -106.6715
22.7662 | -106.6754
22.7698 | -106.6789
22.7735 | -106.6817
22.7771 | -106.6852
22.7808 | -106.6898
22.7845 | -106.693
22.7881 | -106.6931
22.7918 | -106.6922
22.7954 | -106.69
22.7991 | -106.6879
22.8028 | -106.686
22.8064 | -106.6832
22.8101 | -106.6808
22.8137 | -106.6796
22.8174 | -106.6782
22.8211 | -106.6748
22.8247 | -106.669
22.8284 | -106.6648
22.832 | -106.6616
22.8357 | -106.6583
22.8394 | -106.6549
22.843 | -106.6518
22.8467 | -106.6479
22.8503 | -106.6462
22.854 | -106.6468
22.8577 | -106.6423
22.8613 | -106.6374
22.865 | -106.6342
22.8686 | -106.631
22.8723 | -106.6287
22.876 | -106.6277
22.8796 | -106.6265
22.8833 | -106.6241
22.8869 | -106.6208
22.8906 | -106.6172
22.8943 | -106.615
22.8979 | -106.6114
22.9016 | -106.6091
22.9052 | -106.6079
22.9089 | -106.607
22.9126 | -106.6057
22.9162 | -106.6046
22.9199 | -106.6024
22.9235 | -106.5996
22.9272 | -106.5974
22.9309 | -106.5952
22.9345 | -106.5944
22.9382 | -106.5951
22.9418 | -106.5954
22.9455 | -106.5963
22.9492 | -106.5953
22.9528 | -106.5939
22.9565 | -106.5935
22.9601 | -106.5954
22.9638 | -106.596
22.9675 | -106.5948
22.9711 | -106.596
22.9748 | -106.5955
22.9785 | -106.594
22.9821 | -106.5935
22.9858 | -106.5949
22.9894 | -106.5967
22.9931 | -106.5979
22.9968 | -106.5987
23.0004 | -106.5998
23.0041 | -106.599
23.0077 | -106.5976
23.0114 | -106.5975
23.0151 | -106.599
23.0187 | -106.6007
23.0224 | -106.602
23.026 | -106.6008
23.0297 | -106.5993
23.0334 | -106.5978
23.037 | -106.5981
23.0407 | -106.5973
23.0443 | -106.598
23.048 | -106.5995
23.0517 | -106.5987
23.0553 | -106.5977
23.059 | -106.5967
23.0626 | -106.5959
23.0663 | -106.5949
23.07 | -106.594
23.0736 | -106.5924
23.0773 | -106.5926
23.0809 | -106.5923
23.0846 | -106.5887
23.0883 | -106.5873
23.0919 | -106.5854
23.0956 | -106.5822
23.0992 | -106.5791
23.1029 | -106.5762
23.1066 | -106.5728
23.1102 | -106.5707
23.1139 | -106.5685
23.1175 | -106.5668
23.1212 | -106.5664
23.1249 | -106.565
23.1285 | -106.5617
23.1322 | -106.5576
23.1358 | -106.5535
23.1395 | -106.5489
23.1432 | -106.5454
23.1468 | -106.5414
23.1505 | -106.5392
23.1541 | -106.5377
23.1578 | -106.5334
23.1615 | -106.5273
23.1651 | -106.5219
23.1688 | -106.5166
23.1724 | -106.5123
23.1761 | -106.5068
23.1798 | -106.5037
23.1834 | -106.5021
23.1871 | -106.4986
23.1907 | -106.493
23.1944 | -106.4886
23.1981 | -106.4834
23.2017 | -106.478
23.2054 | -106.4738
23.209 | -106.4706
23.2127 | -106.4693
23.2164 | -106.4664
23.22 | -106.4626
23.2237 | -106.4608
23.2273 | -106.4566
23.231 | -106.4549
23.2347 | -106.4526
23.2383 | -106.4494
23.242 | -106.4467
23.2456 | -106.444
23.2493 | -106.4414
23.253 | -106.4388
23.2566 | -106.4364
23.2603 | -106.4309
23.2639 | -106.4248
23.2676 | -106.4202
23.2713 | -106.4161
23.2749 | -106.4112
23.2786 | -106.4078
23.2822 | -106.4054
23.2859 | -106.4023
23.2896 | -106.3989
23.2932 | -106.3937
23.2969 | -106.3892
23.3005 | -106.3863
23.3042 | -106.3843
23.3079 | -106.3826
23.3115 | -106.3781
23.3152 | -106.3731
23.3188 | -106.3688
23.3225 | -106.3634
23.3262 | -106.3593
23.3298 | -106.3547
23.3335 | -106.3514
23.3371 | -106.35
23.3408 | -106.3483
23.3445 | -106.3446
23.3481 | -106.3407
23.3518 | -106.3373
23.3554 | -106.3365
23.3591 | -106.3368
23.3628 | -106.3369
23.3664 | -106.3371
23.3701 | -106.3355
23.3737 | -106.3342
23.3774 | -106.3348
23.3811 | -106.3367
23.3847 | -106.3395
23.3884 | -106.3424
23.392 | -106.3467
23.3957 | -106.3513
23.3994 | -106.3557
23.403 | -106.3605
23.4067 | -106.3659
23.4103 | -106.3721
23.414 | -106.379
23.4177 | -106.3859
23.4213 | -106.3932
23.425 | -106.4016
23.4287 | -106.4086
23.4323 | -106.4164
23.436 | -106.4236
23.4396 | -106.4292
23.4433 | -106.4346
23.447 | -106.4416
23.4506 | -106.4489
23.4543 | -106.4557
23.4579 | -106.4631
23.4616 | -106.4712
23.4653 | -106.4791
23.4689 | -106.4867
23.4726 | -106.4928
23.4762 | -106.5015
23.4799 | -106.5087
23.4836 | -106.5117
23.4872 | -106.5156
23.4909 | -106.5222
23.4945 | -106.528
23.4982 | -106.5335
23.5019 | -106.5394
23.5055 | -106.5463
23.5092 | -106.5499
23.5128 | -106.5538
23.5165 | -106.5603
23.5202 | -106.5665
23.5238 | -106.5706
23.5275 | -106.5754
23.5311 | -106.5819
23.5348 | -106.5875
23.5385 | -106.5914
23.5421 | -106.5927
23.5458 | -106.597
23.5494 | -106.6041
23.5531 | -106.6093
23.5568 | -106.6152
23.5604 | -106.6188
23.5641 | -106.6205
23.5677 | -106.6203
23.5714 | -106.6223
23.5751 | -106.6258
23.5787 | -106.6293
23.5824 | -106.6336
23.586 | -106.6383
23.5897 | -106.6427
23.5934 | -106.6488
23.597 | -106.6548
23.6007 | -106.6576
23.6043 | -106.6617
23.608 | -106.6669
23.6117 | -106.6695
23.6153 | -106.6711
23.619 | -106.6747
23.6226 | -106.6775
23.6263 | -106.6808
23.63 | -106.6843
23.6336 | -106.6861
23.6373 | -106.686
23.6409 | -106.6857
23.6446 | -106.6867
23.6483 | -106.6869
23.6519 | -106.6866
23.6556 | -106.6887
23.6592 | -106.6894
23.6629 | -106.6895
23.6666 | -106.6902
23.6702 | -106.6893
23.6739 | -106.687
23.6775 | -106.6856
23.6812 | -106.6862
23.6849 | -106.6867
23.6885 | -106.6844
23.6922 | -106.6799
23.6958 | -106.6757
23.6995 | -106.6716
23.7032 | -106.6705
23.7068 | -106.6676
23.7105 | -106.6633
23.7141 | -106.6592
23.7178 | -106.6542
23.7215 | -106.6491
23.7251 | -106.6425
23.7288 | -106.635
23.7324 | -106.6286
23.7361 | -106.624
23.7398 | -106.6187
23.7434 | -106.6113
23.7471 | -106.6056
23.7507 | -106.5998
23.7544 | -106.5937
23.7581 | -106.5883
23.7617 | -106.5825
23.7654 | -106.5763
23.769 | -106.5693
23.7727 | -106.5625
23.7764 | -106.5564
23.78 | -106.55
23.7837 | -106.5442
23.7873 | -106.539
23.791 | -106.5336
23.7947 | -106.528
23.7983 | -106.5221
23.802 | -106.5153
23.8056 | -106.5085
23.8093 | -106.5046
23.813 | -106.4996
23.8166 | -106.4953
23.8203 | -106.4931
23.8239 | -106.4902
23.8276 | -106.485
23.8313 | -106.4788
23.8349 | -106.4729
23.8386 | -106.467
23.8422 | -106.4625
23.8459 | -106.4576
23.8496 | -106.4518
23.8532 | -106.4452
23.8569 | -106.4407
23.8605 | -106.4377
23.8642 | -106.435
23.8679 | -106.4314
23.8715 | -106.4283
23.8752 | -106.4248
23.8788 | -106.4233
23.8825 | -106.4224
23.8862 | -106.4193
23.8898 | -106.4165
23.8935 | -106.414
23.8971 | -106.412
23.9008 | -106.4113
23.9045 | -106.4144
23.9081 | -106.4148
23.9118 | -106.4137
23.9154 | -106.414
23.9191 | -106.4151
23.9228 | -106.4147
23.9264 | -106.4153
23.9301 | -106.4151
23.9338 | -106.4135
23.9374 | -106.4141
23.9411 | -106.4151
23.9447 | -106.4164
23.9484 | -106.4197
23.9521 | -106.4238
23.9557 | -106.4264
23.9594 | -106.4286
23.963 | -106.4306
23.9667 | -106.4337
23.9704 | -106.4364
23.974 | -106.4402
23.9777 | -106.4449
23.9813 | -106.4498
23.985 | -106.4538
23.9887 | -106.4565
23.9923 | -106.4573
23.996 | -106.4587
23.9996 | -106.4603
24.0033 | -106.4607
24.007 | -106.4627
24.0106 | -106.466
24.0143 | -106.4657
24.0179 | -106.4661
24.0216 | -106.4665
24.0253 | -106.4678
24.0289 | -106.4706
24.0326 | -106.4721
24.0362 | -106.4711
24.0399 | -106.4691
24.0436 | -106.4681
24.0472 | -106.4669
24.0509 | -106.4641
24.0545 | -106.4617
24.0582 | -106.4599
24.0619 | -106.4598
24.0655 | -106.4601
24.0692 | -106.4586
24.0728 | -106.4555
24.0765 | -106.4522
24.0802 | -106.4496
24.0838 | -106.4465
24.0875 | -106.4445
24.0911 | -106.4432
24.0948 | -106.439
24.0985 | -106.4338
24.1021 | -106.4307
24.1058 | -106.4286
24.1094 | -106.4254
24.1131 | -106.4223
24.1168 | -106.4177
24.1204 | -106.4152
24.1241 | -106.4111
24.1277 | -106.4039
24.1314 | -106.3957
24.1351 | -106.3912
24.1387 | -106.3869
24.1424 | -106.3832
24.146 | -106.3791
24.1497 | -106.3723
24.1534 | -106.3659
24.157 | -106.3589
24.1607 | -106.3518
24.1643 | -106.3451
24.168 | -106.3393
24.1717 | -106.3337
24.1753 | -106.3279
24.179 | -106.3211
24.1826 | -106.313
24.1863 | -106.3067
24.19 | -106.3019
24.1936 | -106.296
24.1973 | -106.2885
24.2009 | -106.2816
24.2046 | -106.2746
24.2083 | -106.2675
24.2119 | -106.2609
24.2156 | -106.2551
24.2192 | -106.2488
24.2229 | -106.2417
24.2266 | -106.2367
24.2302 | -106.2319
24.2339 | -106.2279
24.2375 | -106.2251
24.2412 | -106.222
24.2449 | -106.2161
24.2485 | -106.2101
24.2522 | -106.2047
24.2558 | -106.1985
24.2595 | -106.1916
24.2632 | -106.1872
24.2668 | -106.1788
24.2705 | -106.1567
24.2741 | -106.1153
24.2778 | -106.0662
24.2815 | -106.0066
24.2851 | -105.933
24.2888 | -105.7953
24.2924 | -105.7482
24.2961 | -105.7005
24.2998 | -105.6482
24.3034 | -105.5942
24.3071 | -105.5387
24.3107 | -105.4804
24.3144 | -105.4187
24.3181 | -105.3538
24.3217 | -105.2845
24.3254 | -105.2119
24.329 | -105.1365
24.3327 | -105.0593
24.3364 | -104.9802
24.34 | -104.8996
24.3437 | -104.8176
24.3473 | -104.7342
24.351 | -104.6489
24.3547 | -104.5615
24.3583 | -104.4746
24.362 | -104.3862
24.3656 | -104.2983
24.3693 | -104.2109
24.373 | -104.1245
24.3766 | -104.0391
24.3803 | -103.9562
24.3839 | -103.876
24.3876 | -103.7971
24.3913 | -103.7198
24.3949 | -103.6467
24.3986 | -103.5769
24.4023 | -103.5097
24.4059 | -103.4449
24.4096 | -103.385
24.4132 | -103.3299
24.4169 | -103.2799
24.4206 | -103.2338
24.4242 | -103.1925
24.4279 | -103.1565
24.4315 | -103.125
24.4352 | -103.0989
24.4389 | -103.0793
24.4425 | -103.0656
24.4462 | -103.056
24.4498 | -103.0516
24.4535 | -103.0526
24.4572 | -103.0588
24.4608 | -103.071
24.4645 | -103.0885
24.4681 | -103.1114
24.4718 | -103.1403
24.4755 | -103.1749
24.4791 | -103.2129
24.4828 | -103.2553
24.4864 | -103.3019
24.4901 | -103.3538
24.4938 | -103.4111
24.4974 | -103.4718
24.5011 | -103.5363
24.5047 | -103.6044
24.5084 | -103.6758
24.5121 | -103.7502
24.5157 | -103.8275
24.5194 | -103.9072
24.523 | -103.9882
24.5267 | -104.0702
24.5304 | -104.1522
24.534 | -104.2348
24.5377 | -104.3193
24.5413 | -104.4052
24.545 | -104.4894
24.5487 | -104.573
24.5523 | -104.6548
24.556 | -104.7351
24.5596 | -104.815
24.5633 | -104.8845
24.567 | -104.9411
24.5706 | -104.9799
24.5743 | -105.0059
24.5779 | -105.0011
24.5816 | -104.9741
24.5853 | -104.9383
24.5889 | -104.8953
24.5926 | -104.8301
24.5962 | -104.7689
24.5999 | -104.7177
24.6036 | -104.6695
24.6072 | -104.6149
24.6109 | -104.5514
24.6145 | -104.4868
24.6182 | -104.4225
24.6219 | -104.3461
24.6255 | -104.2645
24.6292 | -104.179
24.6328 | -104.0889
24.6365 | -103.9723
24.6402 | -103.8382
24.6438 | -103.7004
24.6475 | -103.5615
24.6511 | -103.4201
24.6548 | -103.278
24.6585 | -103.1354
24.6621 | -102.9938
24.6658 | -102.8534
24.6694 | -102.7149
24.6731 | -102.5779
24.6768 | -102.4437
24.6804 | -102.3124
24.6841 | -102.1843
24.6877 | -102.0593
24.6914 | -101.9375
24.6951 | -101.8198
24.6987 | -101.7071
24.7024 | -101.6
24.706 | -101.4982
24.7097 | -101.4021
24.7134 | -101.311
24.717 | -101.2256
24.7207 | -101.1466
24.7243 | -101.075
24.728 | -101.0111
24.7317 | -100.9538
24.7353 | -100.9034
24.739 | -100.8602
24.7426 | -100.8237
24.7463 | -100.7946
24.75 | -100.7729
24.7536 | -100.7587
24.7573 | -100.7516
24.7609 | -100.7516
24.7646 | -100.7583
24.7683 | -100.7731
24.7719 | -100.7954
24.7756 | -100.8242
24.7792 | -100.8602
24.7829 | -100.9026
24.7866 | -100.9518
24.7902 | -101.0072
24.7939 | -101.0698
24.7975 | -101.1394
24.8012 | -101.2152
24.8049 | -101.2966
24.8085 | -101.3838
24.8122 | -101.4756
24.8158 | -101.5951
24.8195 | -101.6924
24.8232 | -101.7938
24.8268 | -101.8996
24.8305 | -102.0093
24.8341 | -102.123
24.8378 | -102.2399
24.8415 | -102.3588
24.8451 | -102.479
24.8488 | -102.6017
24.8525 | -102.7255
24.8561 | -102.8492
24.8598 | -102.9737
24.8634 | -103.099
24.8671 | -103.2237
24.8708 | -103.3474
24.8744 | -103.4708
24.8781 | -103.5924
24.8817 | -103.7129
24.8854 | -103.8307
24.8891 | -103.9443
24.8927 | -104.0533
24.8964 | -104.1584
24.9 | -104.2607
24.9037 | -104.3605
24.9074 | -104.4557
24.911 | -104.5456
24.9147 | -104.638
24.9183 | -104.7333
24.922 | -104.8332
24.9257 | -104.9314
24.9293 | -105.0414
24.933 | -105.1524
24.9366 | -105.2498
24.9403 | -105.3329
24.944 | -105.4178
24.9476 | -105.4767
24.9513 | -105.5061
24.9549 | -105.5182
24.9586 | -105.5238
24.9623 | -105.5307
24.9659 | -105.5373
24.9696 | -105.5429
24.9732 | -105.5522
24.9769 | -105.5569
24.9806 | -105.5584
24.9842 | -105.5576
24.9879 | -105.5573
24.9915 | -105.5566
24.9952 | -105.5561
24.9989 | -105.557
25.0025 | -105.5581
25.0062 | -105.5581
25.0098 | -105.5586
25.0135 | -105.5585
25.0172 | -105.5577
25.0208 | -105.5579
25.0245 | -105.5591
25.0281 | -105.5599
25.0318 | -105.5619
25.0355 | -105.5636
25.0391 | -105.5665
25.0428 | -105.5709
25.0464 | -105.5755
25.0501 | -105.58
25.0538 | -105.5849
25.0574 | -105.5903
25.0611 | -105.5952
25.0647 | -105.6022
25.0684 | -105.6109
25.0721 | -105.619
25.0757 | -105.626
25.0794 | -105.6299
25.083 | -105.6354
25.0867 | -105.641
25.0904 | -105.6459
25.094 | -105.6529
25.0977 | -105.6587
25.1013 | -105.6671
25.105 | -105.6753
25.1087 | -105.6841
25.1123 | -105.6937
25.116 | -105.7035
25.1196 | -105.7115
25.1233 | -105.7202
25.127 | -105.7308
25.1306 | -105.741
25.1343 | -105.7528
25.1379 | -105.764
25.1416 | -105.7762
25.1453 | -105.7862
25.1489 | -105.7934
25.1526 | -105.8008
25.1562 | -105.8095
25.1599 | -105.8193
25.1636 | -105.8306
25.1672 | -105.8409
25.1709 | -105.8484
25.1745 | -105.8587
25.1782 | -105.8689
25.1819 | -105.8765
25.1855 | -105.8856
25.1892 | -105.894
25.1928 | -105.9019
25.1965 | -105.9094
25.2002 | -105.9179
25.2038 | -105.9248
25.2075 | -105.9312
25.2111 | -105.9404
25.2148 | -105.9463
25.2185 | -105.9484
25.2221 | -105.9527
25.2258 | -105.9575
25.2294 | -105.961
25.2331 | -105.9654
25.2368 | -105.968
25.2404 | -105.9712
25.2441 | -105.9756
25.2477 | -105.979
25.2514 | -105.9837
25.2551 | -105.9869
25.2587 | -105.9883
25.2624 | -105.99
25.266 | -105.9905
25.2697 | -105.9922
25.2734 | -105.9914
25.277 | -105.9905
25.2807 | -105.9904
25.2843 | -105.9917
25.288 | -105.9932
25.2917 | -105.9952
25.2953 | -105.9991
25.299 | -106.0023
25.3026 | -106.0018
25.3063 | -106.0011
25.31 | -105.9994
25.3136 | -105.9979
25.3173 | -105.9976
25.3209 | -105.9962
25.3246 | -105.995
25.3283 | -105.9914
25.3319 | -105.9904
25.3356 | -105.9906
25.3393 | -105.9908
25.3429 | -105.989
25.3466 | -105.9882
25.3502 | -105.988
25.3539 | -105.9859
25.3576 | -105.9835
25.3612 | -105.9813
25.3649 | -105.9781
25.3685 | -105.9748
25.3722 | -105.9731
25.3759 | -105.9715
25.3795 | -105.9681
25.3832 | -105.9647
25.3868 | -105.9617
25.3905 | -105.9563
25.3942 | -105.9527
25.3978 | -105.9494
25.4015 | -105.9453
25.4051 | -105.9414
25.4088 | -105.9367
25.4125 | -105.9315
25.4161 | -105.9289
25.4198 | -105.925
25.4234 | -105.9189
25.4271 | -105.915
25.4308 | -105.9108
25.4344 | -105.9048
25.4381 | -105.8995
25.4417 | -105.8932
25.4454 | -105.8887
25.4491 | -105.886
25.4527 | -105.8805
25.4564 | -105.8762
25.46 | -105.8698
25.4637 | -105.8646
25.4674 | -105.8594
25.471 | -105.8538
25.4747 | -105.8493
25.4783 | -105.8451
25.482 | -105.8408
25.4857 | -105.8357
25.4893 | -105.8305
25.493 | -105.8261
25.4966 | -105.8222
25.5003 | -105.8182
25.504 | -105.8153
25.5076 | -105.8135
25.5113 | -105.8104
25.5149 | -105.8068
25.5186 | -105.8036
25.5223 | -105.7998
25.5259 | -105.7967
25.5296 | -105.7966
25.5332 | -105.7944
25.5369 | -105.7921
25.5406 | -105.7924
25.5442 | -105.7906
25.5479 | -105.7889
25.5515 | -105.7889
25.5552 | -105.7896
25.5589 | -105.7901
25.5625 | -105.7916
25.5662 | -105.7921
25.5698 | -105.7911
25.5735 | -105.7916
25.5772 | -105.7936
25.5808 | -105.7959
25.5845 | -105.7983
25.5881 | -105.8014
25.5918 | -105.8035
25.5955 | -105.806
25.5991 | -105.8077
25.6028 | -105.8087
25.6064 | -105.8118
25.6101 | -105.8157
25.6138 | -105.8194
25.6174 | -105.823
25.6211 | -105.8272
25.6247 | -105.8316
25.6284 | -105.8352
25.6321 | -105.8399
25.6357 | -105.843
25.6394 | -105.8455
25.643 | -105.8489
25.6467 | -105.8521
25.6504 | -105.8543
25.654 | -105.8549
25.6577 | -105.8571
25.6613 | -105.8614
25.665 | -105.8639
25.6687 | -105.8675
25.6723 | -105.8727
25.676 | -105.8762
25.6796 | -105.8803
25.6833 | -105.8843
25.687 | -105.8887
25.6906 | -105.8912
25.6943 | -105.8938
25.6979 | -105.8969
25.7016 | -105.8996
25.7053 | -105.9015
25.7089 | -105.9043
25.7126 | -105.9062
25.7162 | -105.9085
25.7199 | -105.9108
25.7236 | -105.9137
25.7272 | -105.9179
25.7309 | -105.9215
25.7345 | -105.9245
25.7382 | -105.9276
25.7419 | -105.9316
25.7455 | -105.9365
25.7492 | -105.9398
25.7528 | -105.9436
25.7565 | -105.9452
25.7602 | -105.9475
25.7638 | -105.9506
25.7675 | -105.9527
25.7711 | -105.9568
25.7748 | -105.9596
25.7785 | -105.9614
25.7821 | -105.9653
25.7858 | -105.97
25.7894 | -105.9722
25.7931 | -105.9747
25.7968 | -105.9797
25.8004 | -105.9838
25.8041 | -105.986
25.8077 | -105.9899
25.8114 | -105.9947
25.8151 | -106.0016
25.8187 | -106.008
25.8224 | -106.0141
25.8261 | -106.0191
25.8297 | -106.0235
25.8334 | -106.0281
25.837 | -106.033
25.8407 | -106.0389
25.8444 | -106.0442
25.848 | -106.0489
25.8517 | -106.056
25.8553 | -106.0612
25.859 | -106.0649
25.8627 | -106.0692
25.8663 | -106.0736
25.87 | -106.0794
25.8736 | -106.0877
25.8773 | -106.0951
25.881 | -106.1014
25.8846 | -106.1062
25.8883 | -106.1109
25.8919 | -106.1168
25.8956 | -106.1223
25.8993 | -106.1274
25.9029 | -106.1328
25.9066 | -106.1377
25.9102 | -106.1403
25.9139 | -106.1443
25.9176 | -106.1484
25.9212 | -106.1529
25.9249 | -106.1582
25.9285 | -106.163
25.9322 | -106.1696
25.9359 | -106.1756
25.9395 | -106.1809
25.9432 | -106.1859
25.9468 | -106.1873
25.9505 | -106.1886
25.9542 | -106.1906
25.9578 | -106.195
25.9615 | -106.1989
25.9651 | -106.2031
25.9688 | -106.2073
25.9725 | -106.2101
25.9761 | -106.2126
25.9798 | -106.2158
25.9834 | -106.2186
25.9871 | -106.2197
25.9908 | -106.2211
25.9944 | -106.2214
25.9981 | -106.2213
26.0017 | -106.2232
26.0054 | -106.2251
26.0091 | -106.2253
26.0127 | -106.2251
26.0164 | -106.2248
26.02 | -106.2255
26.0237 | -106.2266
26.0274 | -106.2274
26.031 | -106.2281
26.0347 | -106.2284
26.0383 | -106.2264
26.042 | -106.2247
26.0457 | -106.224
26.0493 | -106.2224
26.053 | -106.2216
26.0566 | -106.2201
26.0603 | -106.2183
26.064 | -106.2175
26.0676 | -106.218
26.0713 | -106.2189
26.0749 | -106.219
26.0786 | -106.2192
26.0823 | -106.2206
26.0859 | -106.2214
26.0896 | -106.2228
26.0932 | -106.2251
26.0969 | -106.2273
26.1006 | -106.228
26.1042 | -106.2302
26.1079 | -106.2339
26.1115 | -106.2365
26.1152 | -106.2408
26.1189 | -106.2448
26.1225 | -106.2489
26.1262 | -106.2539
26.1298 | -106.2588
26.1335 | -106.2656
26.1372 | -106.2707
26.1408 | -106.2764
26.1445 | -106.2823
26.1481 | -106.2881
26.1518 | -106.2925
26.1555 | -106.2971
26.1591 | -106.3035
26.1628 | -106.3079
26.1664 | -106.3122
26.1701 | -106.318
26.1738 | -106.3232
26.1774 | -106.3285
26.1811 | -106.3316
26.1847 | -106.334
26.1884 | -106.3371
26.1921 | -106.3402
26.1957 | -106.3447
26.1994 | -106.3482
26.203 | -106.3511
26.2067 | -106.3543
26.2104 | -106.3571
26.214 | -106.3592
26.2177 | -106.3595
26.2213 | -106.3576
26.225 | -106.3558
26.2287 | -106.3522
26.2323 | -106.3452
26.236 | -106.3329
26.2396 | -106.3195
26.2433 | -106.3077
26.247 | -106.2963
26.2506 | -106.286
26.2543 | -106.2748
26.2579 | -106.2646
26.2616 | -106.2523
26.2653 | -106.2377
26.2689 | -106.2247
26.2726 | -106.2113
26.2763 | -106.197
26.2799 | -106.1831
26.2836 | -106.1664
26.2872 | -106.148
26.2909 | -106.1298
26.2946 | -106.1115
26.2982 | -106.0927
26.3019 | -106.0724
26.3055 | -106.0508
26.3092 | -106.0299
26.3129 | -106.0081
26.3165 | -105.9871
26.3202 | -105.9649
26.3238 | -105.9422
26.3275 | -105.9182
26.3312 | -105.8951
26.3348 | -105.8723
26.3385 | -105.8485
26.3421 | -105.8265
26.3458 | -105.8035
26.3495 | -105.7809
26.3531 | -105.7591
26.3568 | -105.7384
26.3604 | -105.7188
26.3641 | -105.7002
26.3678 | -105.6821
26.3714 | -105.6636
26.3751 | -105.6465
26.3787 | -105.6298
26.3824 | -105.6142
26.3861 | -105.6013
26.3897 | -105.59
26.3934 | -105.5788
26.397 | -105.5692
26.4007 | -105.5768
26.4044 | -105.571
26.408 | -105.5664
26.4117 | -105.5629
26.4153 | -105.5614
26.419 | -105.5615
26.4227 | -105.564
26.4263 | -105.5673
26.43 | -105.5719
26.4336 | -105.5778
26.4373 | -105.5865
26.441 | -105.5961
26.4446 | -105.6077
26.4483 | -105.6204
26.4519 | -105.6342
26.4556 | -105.6494
26.4593 | -105.6665
26.4629 | -105.6831
26.4666 | -105.6994
26.4702 | -105.7183
26.4739 | -105.7388
26.4776 | -105.7601
26.4812 | -105.7809
26.4849 | -105.8024
26.4885 | -105.8238
26.4922 | -105.8449
26.4959 | -105.8668
26.4995 | -105.8877
26.5032 | -105.9094
26.5068 | -105.931
26.5105 | -105.9521
26.5142 | -105.9733
26.5178 | -105.9943
26.5215 | -106.0146
26.5251 | -106.0363
26.5288 | -106.0568
26.5325 | -106.0753
26.5361 | -106.0935
26.5398 | -106.1123
26.5434 | -106.1296
26.5471 | -106.1446
26.5508 | -106.1577
26.5544 | -106.1714
26.5581 | -106.1846
26.5617 | -106.1967
26.5654 | -106.2083
26.5691 | -106.2171
26.5727 | -106.2245
26.5764 | -106.2321
26.58 | -106.236
26.5837 | -106.2394
26.5874 | -106.2449
26.591 | -106.2487
26.5947 | -106.2533
26.5983 | -106.2571
26.602 | -106.2607
26.6057 | -106.2635
26.6093 | -106.2673
26.613 | -106.2685
26.6166 | -106.2671
26.6203 | -106.2638
26.624 | -106.2579
26.6276 | -106.2503
26.6313 | -106.2407
26.6349 | -106.2305
26.6386 | -106.2227
26.6423 | -106.2134
26.6459 | -106.2026
26.6496 | -106.1916
26.6532 | -106.1818
26.6569 | -106.1716
26.6606 | -106.1616
26.6642 | -106.1515
26.6679 | -106.1384
26.6715 | -106.1246
26.6752 | -106.1112
26.6789 | -106.0991
26.6825 | -106.0834
26.6862 | -106.0581
26.6898 | -106.023
26.6935 | -105.9787
26.6972 | -105.9336
26.7008 | -105.8869
26.7045 | -105.8404
26.7081 | -105.8004
26.7118 | -105.7674
26.7155 | -105.7337
26.7191 | -105.6993
26.7228 | -105.6633
26.7264 | -105.6258
26.7301 | -105.5876
26.7338 | -105.5482
26.7374 | -105.5052
26.7411 | -105.4606
26.7447 | -105.4142
26.7484 | -105.3659
26.7521 | -105.3167
26.7557 | -105.2659
26.7594 | -105.2127
26.7631 | -105.1588
26.7667 | -105.1039
26.7704 | -105.0474
26.774 | -104.9899
26.7777 | -104.9325
26.7814 | -104.8735
26.785 | -104.8137
26.7887 | -104.7549
26.7923 | -104.6956
26.796 | -104.6363
26.7997 | -104.5769
26.8033 | -104.5184
26.807 | -104.4609
26.8106 | -104.4047
26.8143 | -104.3482
26.818 | -104.2931
26.8216 | -104.2396
26.8253 | -104.1892
26.8289 | -104.1412
26.8326 | -104.0948
26.8363 | -104.0504
26.8399 | -104.0091
26.8436 | -103.9709
26.8472 | -103.9347
26.8509 | -103.9011
26.8546 | -103.8705
26.8582 | -103.8432
26.8619 | -103.8196
26.8655 | -103.7999
26.8692 | -103.7845
26.8729 | -103.772
26.8765 | -103.7625
26.8802 | -103.7565
26.8838 | -103.7534
26.8875 | -103.7547
26.8912 | -103.7595
26.8948 | -103.7678
26.8985 | -103.7805
26.9021 | -103.7965
26.9058 | -103.8149
26.9095 | -103.8369
26.9131 | -103.8628
26.9168 | -103.891
26.9204 | -103.9231
26.9241 | -103.9585
26.9278 | -104.0179
26.9314 | -104.0563
26.9351 | -104.0971
26.9387 | -104.1401
26.9424 | -104.1854
26.9461 | -104.2319
26.9497 | -104.2806
26.9534 | -104.3315
26.957 | -104.3823
26.9607 | -104.4358
26.9644 | -104.4902
26.968 | -104.5451
26.9717 | -104.6008
26.9753 | -104.6565
26.979 | -104.7125
26.9827 | -104.7686
26.9863 | -104.8255
26.99 | -104.8811
26.9936 | -104.936
26.9973 | -104.9898
27.001 | -105.0433
27.0046 | -105.0953
27.0083 | -105.146
27.0119 | -105.1955
27.0156 | -105.2448
27.0193 | -105.2913
27.0229 | -105.336
27.0266 | -105.3803
27.0302 | -105.4221
27.0339 | -105.4621
27.0376 | -105.4995
27.0412 | -105.5348
27.0449 | -105.5677
27.0485 | -105.6018
27.0522 | -105.6337
27.0559 | -105.6635
27.0595 | -105.6912
27.0632 | -105.7185
27.0668 | -105.7514
27.0705 | -105.7899
27.0742 | -105.8323
27.0778 | -105.8687
27.0815 | -105.9001
27.0851 | -105.9264
27.0888 | -105.9418
27.0925 | -105.9474
27.0961 | -105.9502
27.0998 | -105.9523
27.1034 | -105.9534
27.1071 | -105.9551
27.1108 | -105.9558
27.1144 | -105.9544
27.1181 | -105.956
27.1217 | -105.9581
27.1254 | -105.9602
27.1291 | -105.9617
27.1327 | -105.9643
27.1364 | -105.9652
27.14 | -105.9646
27.1437 | -105.9618
27.1474 | -105.9613
27.151 | -105.9606
27.1547 | -105.9604
27.1583 | -105.9605
27.162 | -105.9604
27.1657 | -105.9606
27.1693 | -105.9577
27.173 | -105.9551
27.1766 | -105.9525
27.1803 | -105.9506
27.184 | -105.9472
27.1876 | -105.9441
27.1913 | -105.9396
27.1949 | -105.9349
27.1986 | -105.9316
27.2023 | -105.9275
27.2059 | -105.9213
27.2096 | -105.9138
27.2132 | -105.9075
27.2169 | -105.8997
27.2206 | -105.894
27.2242 | -105.8883
27.2279 | -105.8828
27.2316 | -105.8765
27.2352 | -105.87
27.2389 | -105.8638
27.2425 | -105.8565
27.2462 | -105.8486
27.2499 | -105.8405
27.2535 | -105.8313
27.2572 | -105.821
27.2608 | -105.8097
27.2645 | -105.7995
27.2682 | -105.7887
27.2718 | -105.7789
27.2755 | -105.7676
27.2791 | -105.7554
27.2828 | -105.7442
27.2865 | -105.7347
27.2901 | -105.7234
27.2938 | -105.7128
27.2974 | -105.7019
27.3011 | -105.6899
27.3048 | -105.6783
27.3084 | -105.6665
27.3121 | -105.6549
27.3157 | -105.6424
27.3194 | -105.6309
27.3231 | -105.6201
27.3267 | -105.6096
27.3304 | -105.5991
27.334 | -105.5881
27.3377 | -105.5768
27.3414 | -105.5673
27.345 | -105.5558
27.3487 | -105.5434
27.3523 | -105.5322
27.356 | -105.5208
27.3597 | -105.5099
27.3633 | -105.4997
27.367 | -105.4894
27.3706 | -105.4796
27.3743 | -105.4715
27.378 | -105.4637
27.3816 | -105.4543
27.3853 | -105.4455
27.3889 | -105.4385
27.3926 | -105.4315
27.3963 | -105.4252
27.3999 | -105.418
27.4036 | -105.4107
27.4072 | -105.406
27.4109 | -105.4006
27.4146 | -105.3964
27.4182 | -105.3931
27.4219 | -105.3901
27.4255 | -105.3865
27.4292 | -105.3836
27.4329 | -105.3802
27.4365 | -105.379
27.4402 | -105.3795
27.4438 | -105.3802
27.4475 | -105.3813
27.4512 | -105.382
27.4548 | -105.3877
27.4585 | -105.3889
27.4621 | -105.3898
27.4658 | -105.3915
27.4695 | -105.3936
27.4731 | -105.3961
27.4768 | -105.3981
27.4804 | -105.4017
27.4841 | -105.4063
27.4878 | -105.4123
27.4914 | -105.4184
27.4951 | -105.4242
27.4987 | -105.4299
27.5024 | -105.4358
27.5061 | -105.4416
27.5097 | -105.4471
27.5134 | -105.4522
27.517 | -105.4578
27.5207 | -105.463
27.5244 | -105.4689
27.528 | -105.4766
27.5317 | -105.485
27.5353 | -105.4913
27.539 | -105.4971
27.5427 | -105.501
27.5463 | -105.5058
27.55 | -105.512
27.5536 | -105.5171
27.5573 | -105.5222
27.561 | -105.5268
27.5646 | -105.5314
27.5683 | -105.5349
27.5719 | -105.5389
27.5756 | -105.5423
27.5793 | -105.5462
27.5829 | -105.5496
27.5866 | -105.5518
27.5902 | -105.5538
27.5939 | -105.5558
27.5976 | -105.558
27.6012 | -105.559
27.6049 | -105.5595
27.6085 | -105.5575
27.6122 | -105.5554
27.6159 | -105.554
27.6195 | -105.5531
27.6232 | -105.5506
27.6268 | -105.5471
27.6305 | -105.544
27.6342 | -105.5396
27.6378 | -105.5362
27.6415 | -105.5321
27.6451 | -105.5294
27.6488 | -105.5259
27.6525 | -105.5206
27.6561 | -105.5148
27.6598 | -105.5081
27.6634 | -105.5021
27.6671 | -105.495
27.6708 | -105.4863
27.6744 | -105.4769
27.6781 | -105.4695
27.6818 | -105.4599
27.6854 | -105.4509
27.6891 | -105.4428
27.6927 | -105.4343
27.6964 | -105.4257
27.7001 | -105.4167
27.7037 | -105.408
27.7074 | -105.3988
27.711 | -105.3899
27.7147 | -105.3816
27.7184 | -105.3731
27.722 | -105.3637
27.7257 | -105.3546
27.7293 | -105.3457
27.733 | -105.3376
27.7367 | -105.3295
27.7403 | -105.3208
27.744 | -105.3123
27.7476 | -105.3039
27.7513 | -105.2952
27.755 | -105.2863
27.7586 | -105.2785
27.7623 | -105.2724
27.7659 | -105.2661
27.7696 | -105.2609
27.7733 | -105.2546
27.7769 | -105.2492
27.7806 | -105.2444
27.7842 | -105.2376
27.7879 | -105.2314
27.7916 | -105.226
27.7952 | -105.2207
27.7989 | -105.2165
27.8025 | -105.2112
27.8062 | -105.2059
27.8099 | -105.2023
27.8135 | -105.1992
27.8172 | -105.1947
27.8208 | -105.1911
27.8245 | -105.1878
27.8282 | -105.1842
27.8318 | -105.181
27.8355 | -105.1772
27.8391 | -105.1747
27.8428 | -105.1729
27.8465 | -105.1716
27.8501 | -105.1684
27.8538 | -105.166
27.8574 | -105.1636
27.8611 | -105.1611
27.8648 | -105.1571
27.8684 | -105.1553
27.8721 | -105.1532
27.8757 | -105.1507
27.8794 | -105.1483
27.8831 | -105.1473
27.8867 | -105.1453
27.8904 | -105.1423
27.894 | -105.1423
27.8977 | -105.1418
27.9014 | -105.1416
27.905 | -105.14
27.9087 | -105.1391
27.9123 | -105.1374
27.916 | -105.1372
27.9197 | -105.1378
27.9233 | -105.1358
27.927 | -105.1295
27.9306 | -105.1217
27.9343 | -105.1155
27.938 | -105.1108
27.9416 | -105.1082
27.9453 | -105.1057
27.9489 | -105.1043
27.9526 | -105.1027
27.9563 | -105.1008
27.9599 | -105.0987
27.9636 | -105.0966
27.9672 | -105.0933
27.9709 | -105.0899
27.9746 | -105.0881
27.9782 | -105.0846
27.9819 | -105.0818
27.9855 | -105.0805
27.9892 | -105.0786
27.9929 | -105.0758
27.9965 | -105.0709
28.0002 | -105.0681