## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 282 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 285
Hood | 02-92B/P | 185
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1011.2 | hPa |  | PRC PRO 01
Humidity | 54.4 | %/Hr
Internal temperature of calibration bench | 24.13 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1020 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.8 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.95 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.99 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3103 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 03/09/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 07/09/2018
Technical Manager | Anthony HUE | Approval |  | 27/09/2018
CEO | Robert GLEMOT | Approval |  | 27/09/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3865 | 57.3303
0.012209 | -35.9823 | 38.1116
0.018313 | -34.9321 | 27.9841
0.024417 | -34.4789 | 21.7426
0.030521 | -34.3685 | 17.9939
0.036626 | -34.2279 | 14.9525
0.04273 | -34.136 | 12.6048
0.048834 | -34.0921 | 11.082
0.054938 | -34.0299 | 9.9446
0.061043 | -33.9716 | 8.8703
0.067147 | -33.9623 | 8.0549
0.073251 | -33.9639 | 7.2982
0.079355 | -33.9821 | 6.6389
0.08546 | -33.9961 | 6.0499
0.091564 | -33.992 | 5.6125
0.097668 | -33.9753 | 5.3607
0.10377 | -33.9549 | 5.1111
0.10988 | -33.9391 | 4.8336
0.11598 | -33.9308 | 4.5427
0.12209 | -33.9331 | 4.2617
0.12819 | -33.9432 | 3.9824
0.13429 | -33.9543 | 3.6996
0.1404 | -33.9579 | 3.4267
0.1465 | -33.9602 | 3.1736
0.15261 | -33.9578 | 2.9569
0.15871 | -33.9518 | 2.7962
0.16482 | -33.9444 | 2.6928
0.17092 | -33.9428 | 2.6508
0.17702 | -33.9413 | 2.5771
0.18313 | -33.941 | 2.4978
0.18923 | -33.9412 | 2.3892
0.19534 | -33.9404 | 2.2637
0.20144 | -33.9395 | 2.1234
0.20754 | -33.937 | 2.0038
0.21365 | -33.9363 | 1.8904
0.21975 | -33.9362 | 1.7932
0.22586 | -33.9376 | 1.7064
0.23196 | -33.9403 | 1.6345
0.23807 | -33.9445 | 1.5722
0.24417 | -33.9488 | 1.5141
0.25027 | -33.9529 | 1.4522
0.25638 | -33.9554 | 1.3909
0.26248 | -33.9575 | 1.3251
0.26859 | -33.9587 | 1.2607
0.27469 | -33.9593 | 1.1908
0.2808 | -33.9606 | 1.1185
0.2869 | -33.9617 | 1.0488
0.293 | -33.9628 | 0.98288
0.29911 | -33.964 | 0.9238
0.30521 | -33.9652 | 0.86563
0.31132 | -33.9661 | 0.8272
0.31742 | -33.9675 | 0.78473
0.32353 | -33.968 | 0.74971
0.32963 | -33.969 | 0.72046
0.33573 | -33.9701 | 0.69949
0.34184 | -33.9715 | 0.67751
0.34794 | -33.9737 | 0.65319
0.35405 | -33.9759 | 0.62862
0.36015 | -33.9785 | 0.60288
0.36626 | -33.9805 | 0.57509
0.37236 | -33.9827 | 0.53859
0.37846 | -33.9849 | 0.49488
0.38457 | -33.986 | 0.43922
0.39067 | -33.9871 | 0.39347
0.39678 | -33.9879 | 0.34281
0.40288 | -33.9886 | 0.29015
0.40899 | -33.9887 | 0.23963
0.41509 | -33.9882 | 0.19015
0.42119 | -33.9884 | 0.14799
0.4273 | -33.9886 | 0.10297
0.4334 | -33.9888 | 0.057954
0.43951 | -33.9893 | 0.015511
0.44561 | -33.9903 | -0.023486
0.45172 | -33.992 | -0.054991
0.45782 | -33.9929 | -0.091566
0.46392 | -33.9943 | -0.12728
0.47003 | -33.9955 | -0.15944
0.47613 | -33.9965 | -0.19324
0.48224 | -33.9972 | -0.22777
0.48834 | -33.9964 | -0.26241
0.49445 | -33.9963 | -0.29744
0.50055 | -33.9959 | -0.32947
0.50665 | -33.9954 | -0.36354
0.51276 | -33.9945 | -0.39747
0.51886 | -33.9936 | -0.43124
0.52497 | -33.9926 | -0.45551
0.53107 | -33.992 | -0.48493
0.53717 | -33.9914 | -0.51303
0.54328 | -33.9907 | -0.53602
0.54938 | -33.9907 | -0.55856
0.55549 | -33.9908 | -0.57962
0.56159 | -33.992 | -0.59158
0.5677 | -33.9924 | -0.60641
0.5738 | -33.9936 | -0.62294
0.5799 | -33.9945 | -0.63809
0.58601 | -33.9956 | -0.64859
0.59211 | -33.9973 | -0.65546
0.59822 | -33.9988 | -0.66537
0.60432 | -34.0003 | -0.67778
0.61043 | -34.0017 | -0.68616
0.61653 | -34.0029 | -0.69679
0.62263 | -34.0045 | -0.70652
0.62874 | -34.0053 | -0.71735
0.63484 | -34.0062 | -0.7269
0.64095 | -34.0067 | -0.73732
0.64705 | -34.0072 | -0.74731
0.65316 | -34.0075 | -0.75991
0.65926 | -34.0076 | -0.77576
0.66536 | -34.0067 | -0.79308
0.67147 | -34.0064 | -0.80636
0.67757 | -34.0061 | -0.82069
0.68368 | -34.0055 | -0.83943
0.68978 | -34.0048 | -0.85647
0.69589 | -34.004 | -0.87532
0.70199 | -34.0032 | -0.89776
0.70809 | -34.0025 | -0.91672
0.7142 | -34.0019 | -0.93696
0.7203 | -34.001 | -0.95597
0.72641 | -34.0005 | -0.97821
0.73251 | -34 | -0.99804
0.73862 | -34.0004 | -1.0195
0.74472 | -34.0002 | -1.0394
0.75082 | -34.0001 | -1.0612
0.75693 | -34.0005 | -1.0828
0.76303 | -34.0011 | -1.1046
0.76914 | -34.003 | -1.1248
0.77524 | -34.0041 | -1.1443
0.78135 | -34.0054 | -1.1642
0.78745 | -34.0068 | -1.184
0.79355 | -34.0081 | -1.2036
0.79966 | -34.0099 | -1.2229
0.80576 | -34.0122 | -1.2424
0.81187 | -34.0138 | -1.2638
0.81797 | -34.0158 | -1.2849
0.82408 | -34.0175 | -1.3052
0.83018 | -34.0193 | -1.3256
0.83628 | -34.0212 | -1.344
0.84239 | -34.0224 | -1.3644
0.84849 | -34.0238 | -1.3859
0.8546 | -34.0253 | -1.4055
0.8607 | -34.0266 | -1.4261
0.86681 | -34.0277 | -1.445
0.87291 | -34.0287 | -1.4637
0.87901 | -34.0292 | -1.4816
0.88512 | -34.03 | -1.4997
0.89122 | -34.0307 | -1.5179
0.89733 | -34.0312 | -1.5335
0.90343 | -34.0317 | -1.551
0.90953 | -34.0319 | -1.5653
0.91564 | -34.032 | -1.5811
0.92174 | -34.0323 | -1.5961
0.92785 | -34.0324 | -1.6107
0.93395 | -34.0326 | -1.6245
0.94006 | -34.0328 | -1.6373
0.94616 | -34.0326 | -1.6491
0.95226 | -34.0328 | -1.6617
0.95837 | -34.033 | -1.6755
0.96447 | -34.0332 | -1.6883
0.97058 | -34.0332 | -1.7016
0.97668 | -34.033 | -1.7148
0.98279 | -34.0328 | -1.7274
0.98889 | -34.0327 | -1.738
0.99499 | -34.0328 | -1.7513
1.0011 | -34.0327 | -1.762
1.0072 | -34.0325 | -1.7736
1.0133 | -34.0322 | -1.7843
1.0194 | -34.0319 | -1.7972
1.0255 | -34.0316 | -1.8107
1.0316 | -34.0312 | -1.8225
1.0377 | -34.0309 | -1.8361
1.0438 | -34.0304 | -1.8486
1.0499 | -34.0297 | -1.8637
1.056 | -34.0291 | -1.8779
1.0621 | -34.0286 | -1.892
1.0682 | -34.0282 | -1.9044
1.0743 | -34.0276 | -1.9197
1.0805 | -34.0272 | -1.9339
1.0866 | -34.0264 | -1.9489
1.0927 | -34.0257 | -1.9633
1.0988 | -34.0251 | -1.978
1.1049 | -34.0244 | -1.9921
1.111 | -34.0237 | -2.0073
1.1171 | -34.0231 | -2.0217
1.1232 | -34.0227 | -2.0368
1.1293 | -34.0224 | -2.0507
1.1354 | -34.0222 | -2.0657
1.1415 | -34.0218 | -2.081
1.1476 | -34.0213 | -2.0964
1.1537 | -34.0209 | -2.1109
1.1598 | -34.021 | -2.1257
1.1659 | -34.0209 | -2.139
1.172 | -34.0207 | -2.1528
1.1781 | -34.0205 | -2.1667
1.1842 | -34.0204 | -2.1807
1.1903 | -34.0208 | -2.1918
1.1964 | -34.0207 | -2.2051
1.2025 | -34.0208 | -2.2189
1.2086 | -34.0212 | -2.2314
1.2147 | -34.0213 | -2.2439
1.2209 | -34.0216 | -2.2551
1.227 | -34.0224 | -2.2659
1.2331 | -34.0228 | -2.2778
1.2392 | -34.0232 | -2.2896
1.2453 | -34.0236 | -2.3007
1.2514 | -34.0241 | -2.3106
1.2575 | -34.0247 | -2.3216
1.2636 | -34.0255 | -2.3304
1.2697 | -34.0261 | -2.3397
1.2758 | -34.0268 | -2.3507
1.2819 | -34.0275 | -2.3607
1.288 | -34.0282 | -2.3714
1.2941 | -34.0291 | -2.3799
1.3002 | -34.0298 | -2.388
1.3063 | -34.0303 | -2.3971
1.3124 | -34.031 | -2.4061
1.3185 | -34.0317 | -2.4151
1.3246 | -34.0325 | -2.4242
1.3307 | -34.0333 | -2.4328
1.3368 | -34.0341 | -2.4415
1.3429 | -34.0348 | -2.4504
1.349 | -34.0353 | -2.4597
1.3551 | -34.036 | -2.469
1.3613 | -34.0366 | -2.4772
1.3674 | -34.0373 | -2.4872
1.3735 | -34.0378 | -2.4969
1.3796 | -34.0383 | -2.506
1.3857 | -34.0386 | -2.5152
1.3918 | -34.0392 | -2.5254
1.3979 | -34.0398 | -2.5346
1.404 | -34.0403 | -2.5459
1.4101 | -34.0408 | -2.5562
1.4162 | -34.0412 | -2.5668
1.4223 | -34.0417 | -2.5781
1.4284 | -34.0422 | -2.5881
1.4345 | -34.0425 | -2.5996
1.4406 | -34.0429 | -2.6131
1.4467 | -34.0434 | -2.6246
1.4528 | -34.0438 | -2.636
1.4589 | -34.0444 | -2.6478
1.465 | -34.0446 | -2.6587
1.4711 | -34.0452 | -2.6701
1.4772 | -34.0457 | -2.6833
1.4833 | -34.0462 | -2.6962
1.4894 | -34.0468 | -2.7084
1.4955 | -34.0473 | -2.7204
1.5016 | -34.0479 | -2.7328
1.5078 | -34.0483 | -2.7467
1.5139 | -34.0488 | -2.7597
1.52 | -34.0492 | -2.7723
1.5261 | -34.0496 | -2.7849
1.5322 | -34.0499 | -2.7989
1.5383 | -34.0503 | -2.8115
1.5444 | -34.0509 | -2.8242
1.5505 | -34.0514 | -2.8375
1.5566 | -34.0519 | -2.8506
1.5627 | -34.0524 | -2.8643
1.5688 | -34.053 | -2.8779
1.5749 | -34.0534 | -2.8906
1.581 | -34.0537 | -2.904
1.5871 | -34.0542 | -2.9177
1.5932 | -34.0546 | -2.9306
1.5993 | -34.055 | -2.9448
1.6054 | -34.0554 | -2.9581
1.6115 | -34.0558 | -2.9722
1.6176 | -34.0561 | -2.9848
1.6237 | -34.0564 | -2.9985
1.6298 | -34.0567 | -3.0111
1.6359 | -34.0571 | -3.024
1.642 | -34.0575 | -3.036
1.6482 | -34.0578 | -3.0491
1.6543 | -34.0581 | -3.0629
1.6604 | -34.0585 | -3.0749
1.6665 | -34.0586 | -3.0875
1.6726 | -34.0589 | -3.1001
1.6787 | -34.0592 | -3.1128
1.6848 | -34.0592 | -3.1255
1.6909 | -34.0595 | -3.1383
1.697 | -34.0597 | -3.1507
1.7031 | -34.0598 | -3.1634
1.7092 | -34.0598 | -3.1757
1.7153 | -34.0601 | -3.1877
1.7214 | -34.0599 | -3.2
1.7275 | -34.0599 | -3.2127
1.7336 | -34.0599 | -3.2256
1.7397 | -34.0601 | -3.238
1.7458 | -34.0601 | -3.2503
1.7519 | -34.0602 | -3.2625
1.758 | -34.0603 | -3.2746
1.7641 | -34.0604 | -3.2862
1.7702 | -34.0604 | -3.2992
1.7763 | -34.0605 | -3.3108
1.7824 | -34.0605 | -3.3233
1.7885 | -34.0605 | -3.3358
1.7947 | -34.0607 | -3.3478
1.8008 | -34.0607 | -3.3591
1.8069 | -34.0608 | -3.3708
1.813 | -34.0609 | -3.383
1.8191 | -34.0609 | -3.3957
1.8252 | -34.061 | -3.4083
1.8313 | -34.0612 | -3.4206
1.8374 | -34.0614 | -3.4324
1.8435 | -34.0615 | -3.4443
1.8496 | -34.0616 | -3.4572
1.8557 | -34.0618 | -3.4699
1.8618 | -34.0621 | -3.482
1.8679 | -34.0623 | -3.4939
1.874 | -34.0625 | -3.5058
1.8801 | -34.0627 | -3.5181
1.8862 | -34.063 | -3.5301
1.8923 | -34.0631 | -3.5421
1.8984 | -34.0633 | -3.554
1.9045 | -34.0634 | -3.5655
1.9106 | -34.0635 | -3.5776
1.9167 | -34.0636 | -3.5897
1.9228 | -34.0638 | -3.6014
1.9289 | -34.064 | -3.6131
1.9351 | -34.0642 | -3.6238
1.9412 | -34.0644 | -3.635
1.9473 | -34.0648 | -3.6469
1.9534 | -34.065 | -3.6582
1.9595 | -34.0652 | -3.6699
1.9656 | -34.0655 | -3.682
1.9717 | -34.0658 | -3.6936
1.9778 | -34.0661 | -3.7049
1.9839 | -34.0664 | -3.7162
1.99 | -34.0666 | -3.7271
1.9961 | -34.0668 | -3.7391
2.0022 | -34.067 | -3.7504
2.0083 | -34.0671 | -3.7615
2.0144 | -34.0674 | -3.7728
2.0205 | -34.0677 | -3.7844
2.0266 | -34.0681 | -3.7957
2.0327 | -34.0684 | -3.8071
2.0388 | -34.0688 | -3.8186
2.0449 | -34.069 | -3.8298
2.051 | -34.0693 | -3.8411
2.0571 | -34.0697 | -3.8524
2.0632 | -34.0699 | -3.8637
2.0693 | -34.0701 | -3.8758
2.0754 | -34.0704 | -3.8871
2.0816 | -34.0707 | -3.8991
2.0877 | -34.0709 | -3.9105
2.0938 | -34.0713 | -3.9225
2.0999 | -34.0715 | -3.9341
2.106 | -34.0717 | -3.9456
2.1121 | -34.0719 | -3.9575
2.1182 | -34.0722 | -3.9682
2.1243 | -34.0724 | -3.979
2.1304 | -34.0726 | -3.99
2.1365 | -34.073 | -4.0017
2.1426 | -34.0732 | -4.0136
2.1487 | -34.0735 | -4.0248
2.1548 | -34.0739 | -4.0365
2.1609 | -34.0743 | -4.0479
2.167 | -34.0747 | -4.059
2.1731 | -34.075 | -4.0709
2.1792 | -34.0752 | -4.082
2.1853 | -34.0756 | -4.0938
2.1914 | -34.0759 | -4.1051
2.1975 | -34.0763 | -4.1165
2.2036 | -34.0765 | -4.1285
2.2097 | -34.0768 | -4.1399
2.2158 | -34.0772 | -4.1505
2.222 | -34.0774 | -4.161
2.2281 | -34.0777 | -4.1723
2.2342 | -34.078 | -4.1836
2.2403 | -34.0782 | -4.1942
2.2464 | -34.0784 | -4.205
2.2525 | -34.0787 | -4.2163
2.2586 | -34.079 | -4.2269
2.2647 | -34.0793 | -4.2373
2.2708 | -34.0796 | -4.2485
2.2769 | -34.08 | -4.2592
2.283 | -34.0802 | -4.2703
2.2891 | -34.0804 | -4.2812
2.2952 | -34.0806 | -4.2917
2.3013 | -34.0808 | -4.3026
2.3074 | -34.0811 | -4.313
2.3135 | -34.0813 | -4.3244
2.3196 | -34.0815 | -4.3345
2.3257 | -34.0817 | -4.3451
2.3318 | -34.0818 | -4.3561
2.3379 | -34.082 | -4.3669
2.344 | -34.0822 | -4.3782
2.3501 | -34.0824 | -4.3889
2.3562 | -34.0826 | -4.3993
2.3623 | -34.0827 | -4.4097
2.3685 | -34.0828 | -4.4203
2.3746 | -34.083 | -4.4315
2.3807 | -34.0832 | -4.442
2.3868 | -34.0834 | -4.4527
2.3929 | -34.0836 | -4.4633
2.399 | -34.0837 | -4.4738
2.4051 | -34.0839 | -4.4843
2.4112 | -34.0841 | -4.4949
2.4173 | -34.0843 | -4.5055
2.4234 | -34.0844 | -4.5163
2.4295 | -34.0846 | -4.5272
2.4356 | -34.0847 | -4.538
2.4417 | -34.0849 | -4.5491
2.4478 | -34.085 | -4.56
2.4539 | -34.0852 | -4.5702
2.46 | -34.0853 | -4.5817
2.4661 | -34.0855 | -4.5928
2.4722 | -34.0857 | -4.603
2.4783 | -34.0858 | -4.6134
2.4844 | -34.0859 | -4.6244
2.4905 | -34.086 | -4.6354
2.4966 | -34.0862 | -4.6461
2.5027 | -34.0863 | -4.6566
2.5089 | -34.0865 | -4.6675
2.515 | -34.0865 | -4.6783
2.5211 | -34.0866 | -4.6891
2.5272 | -34.0867 | -4.7
2.5333 | -34.0869 | -4.711
2.5394 | -34.0871 | -4.7217
2.5455 | -34.0872 | -4.733
2.5516 | -34.0874 | -4.7434
2.5577 | -34.0876 | -4.7549
2.5638 | -34.0878 | -4.7666
2.5699 | -34.0879 | -4.778
2.576 | -34.088 | -4.7896
2.5821 | -34.0881 | -4.8008
2.5882 | -34.0882 | -4.8119
2.5943 | -34.0884 | -4.8228
2.6004 | -34.0886 | -4.8337
2.6065 | -34.0887 | -4.8441
2.6126 | -34.0889 | -4.8547
2.6187 | -34.0891 | -4.8652
2.6248 | -34.0892 | -4.876
2.6309 | -34.0894 | -4.8874
2.637 | -34.0897 | -4.899
2.6431 | -34.0899 | -4.9096
2.6492 | -34.09 | -4.9209
2.6554 | -34.0902 | -4.9318
2.6615 | -34.0903 | -4.9428
2.6676 | -34.0906 | -4.9542
2.6737 | -34.0909 | -4.9654
2.6798 | -34.0912 | -4.9765
2.6859 | -34.0914 | -4.9878
2.692 | -34.0916 | -4.9987
2.6981 | -34.0918 | -5.0098
2.7042 | -34.0921 | -5.021
2.7103 | -34.0923 | -5.0316
2.7164 | -34.0926 | -5.0426
2.7225 | -34.0929 | -5.054
2.7286 | -34.0931 | -5.065
2.7347 | -34.0933 | -5.0756
2.7408 | -34.0936 | -5.0857
2.7469 | -34.0939 | -5.097
2.753 | -34.0942 | -5.1076
2.7591 | -34.0945 | -5.1189
2.7652 | -34.0947 | -5.1304
2.7713 | -34.0949 | -5.1412
2.7774 | -34.0952 | -5.1522
2.7835 | -34.0954 | -5.1632
2.7896 | -34.0956 | -5.1737
2.7958 | -34.0959 | -5.1845
2.8019 | -34.0962 | -5.195
2.808 | -34.0965 | -5.2058
2.8141 | -34.0968 | -5.217
2.8202 | -34.0972 | -5.2279
2.8263 | -34.0975 | -5.2388
2.8324 | -34.0978 | -5.2502
2.8385 | -34.0982 | -5.2611
2.8446 | -34.0986 | -5.2718
2.8507 | -34.099 | -5.2826
2.8568 | -34.0994 | -5.2932
2.8629 | -34.0996 | -5.3038
2.869 | -34.1 | -5.3149
2.8751 | -34.1003 | -5.3256
2.8812 | -34.1007 | -5.3369
2.8873 | -34.101 | -5.3478
2.8934 | -34.1013 | -5.3589
2.8995 | -34.1017 | -5.37
2.9056 | -34.102 | -5.3811
2.9117 | -34.1023 | -5.3921
2.9178 | -34.1027 | -5.4029
2.9239 | -34.1031 | -5.4144
2.93 | -34.1035 | -5.4258
2.9361 | -34.1038 | -5.4367
2.9423 | -34.1041 | -5.4473
2.9484 | -34.1046 | -5.4584
2.9545 | -34.1049 | -5.4694
2.9606 | -34.1053 | -5.4808
2.9667 | -34.1057 | -5.4919
2.9728 | -34.106 | -5.5031
2.9789 | -34.1065 | -5.5142
2.985 | -34.1069 | -5.5254
2.9911 | -34.1073 | -5.5367
2.9972 | -34.1077 | -5.5478
3.0033 | -34.1081 | -5.5592
3.0094 | -34.1085 | -5.5703
3.0155 | -34.1089 | -5.5812
3.0216 | -34.1093 | -5.5924
3.0277 | -34.1096 | -5.6041
3.0338 | -34.11 | -5.6154
3.0399 | -34.1104 | -5.6267
3.046 | -34.1107 | -5.6382
3.0521 | -34.1111 | -5.6492
3.0582 | -34.1114 | -5.6612
3.0643 | -34.1117 | -5.6726
3.0704 | -34.112 | -5.6843
3.0765 | -34.1123 | -5.6954
3.0827 | -34.1126 | -5.7066
3.0888 | -34.1129 | -5.7178
3.0949 | -34.1133 | -5.729
3.101 | -34.1137 | -5.7404
3.1071 | -34.114 | -5.7521
3.1132 | -34.1144 | -5.7638
3.1193 | -34.1147 | -5.7752
3.1254 | -34.1151 | -5.7865
3.1315 | -34.1155 | -5.7983
3.1376 | -34.1158 | -5.8096
3.1437 | -34.1161 | -5.8211
3.1498 | -34.1165 | -5.8331
3.1559 | -34.1168 | -5.8445
3.162 | -34.117 | -5.8564
3.1681 | -34.1173 | -5.8683
3.1742 | -34.1176 | -5.8797
3.1803 | -34.1179 | -5.8913
3.1864 | -34.1183 | -5.9025
3.1925 | -34.1186 | -5.914
3.1986 | -34.1188 | -5.9257
3.2047 | -34.1191 | -5.9377
3.2108 | -34.1194 | -5.9487
3.2169 | -34.1197 | -5.9603
3.223 | -34.1199 | -5.9719
3.2292 | -34.1202 | -5.9837
3.2353 | -34.1204 | -5.9944
3.2414 | -34.1206 | -6.0057
3.2475 | -34.1208 | -6.0171
3.2536 | -34.1211 | -6.0284
3.2597 | -34.1214 | -6.0393
3.2658 | -34.1216 | -6.0502
3.2719 | -34.1218 | -6.0612
3.278 | -34.1221 | -6.0727
3.2841 | -34.1223 | -6.0837
3.2902 | -34.1226 | -6.095
3.2963 | -34.1228 | -6.1064
3.3024 | -34.1229 | -6.1174
3.3085 | -34.1232 | -6.1289
3.3146 | -34.1234 | -6.14
3.3207 | -34.1237 | -6.1515
3.3268 | -34.1239 | -6.1629
3.3329 | -34.1241 | -6.1735
3.339 | -34.1242 | -6.1842
3.3451 | -34.1244 | -6.1954
3.3512 | -34.1246 | -6.206
3.3573 | -34.1249 | -6.2172
3.3634 | -34.1251 | -6.2282
3.3696 | -34.1253 | -6.2395
3.3757 | -34.1254 | -6.2502
3.3818 | -34.1257 | -6.2611
3.3879 | -34.1258 | -6.2721
3.394 | -34.126 | -6.2834
3.4001 | -34.1262 | -6.2942
3.4062 | -34.1263 | -6.3051
3.4123 | -34.1265 | -6.3156
3.4184 | -34.1267 | -6.3268
3.4245 | -34.1269 | -6.3376
3.4306 | -34.127 | -6.3482
3.4367 | -34.1272 | -6.3587
3.4428 | -34.1274 | -6.3691
3.4489 | -34.1275 | -6.38
3.455 | -34.1277 | -6.3908
3.4611 | -34.1279 | -6.402
3.4672 | -34.1281 | -6.413
3.4733 | -34.1282 | -6.4241
3.4794 | -34.1284 | -6.4349
3.4855 | -34.1286 | -6.4453
3.4916 | -34.1288 | -6.4557
3.4977 | -34.129 | -6.4669
3.5038 | -34.1291 | -6.4777
3.5099 | -34.1293 | -6.488
3.5161 | -34.1295 | -6.4984
3.5222 | -34.1296 | -6.5088
3.5283 | -34.1297 | -6.5196
3.5344 | -34.1298 | -6.5304
3.5405 | -34.13 | -6.5409
3.5466 | -34.1302 | -6.5518
3.5527 | -34.1303 | -6.5627
3.5588 | -34.1305 | -6.5733
3.5649 | -34.1305 | -6.584
3.571 | -34.1307 | -6.5946
3.5771 | -34.1308 | -6.6057
3.5832 | -34.131 | -6.6156
3.5893 | -34.1312 | -6.6262
3.5954 | -34.1314 | -6.6371
3.6015 | -34.1316 | -6.648
3.6076 | -34.1318 | -6.6585
3.6137 | -34.132 | -6.6686
3.6198 | -34.1323 | -6.679
3.6259 | -34.1324 | -6.6899
3.632 | -34.1326 | -6.7005
3.6381 | -34.1328 | -6.7113
3.6442 | -34.1329 | -6.7219
3.6503 | -34.1332 | -6.7322
3.6565 | -34.1334 | -6.743
3.6626 | -34.1335 | -6.7535
3.6687 | -34.1338 | -6.7643
3.6748 | -34.1341 | -6.7747
3.6809 | -34.1343 | -6.7854
3.687 | -34.1345 | -6.7964
3.6931 | -34.1348 | -6.8067
3.6992 | -34.135 | -6.8174
3.7053 | -34.1352 | -6.828
3.7114 | -34.1354 | -6.8387
3.7175 | -34.1356 | -6.8494
3.7236 | -34.1359 | -6.8596
3.7297 | -34.1361 | -6.8704
3.7358 | -34.1363 | -6.8813
3.7419 | -34.1365 | -6.8922
3.748 | -34.1368 | -6.9027
3.7541 | -34.137 | -6.9134
3.7602 | -34.1373 | -6.924
3.7663 | -34.1375 | -6.9346
3.7724 | -34.1377 | -6.9453
3.7785 | -34.138 | -6.9558
3.7846 | -34.1382 | -6.9663
3.7907 | -34.1385 | -6.9772
3.7969 | -34.1388 | -6.988
3.803 | -34.139 | -6.9983
3.8091 | -34.1393 | -7.0089
3.8152 | -34.1396 | -7.0195
3.8213 | -34.1398 | -7.0304
3.8274 | -34.1401 | -7.0408
3.8335 | -34.1404 | -7.0516
3.8396 | -34.1407 | -7.0623
3.8457 | -34.141 | -7.0731
3.8518 | -34.1414 | -7.0837
3.8579 | -34.1417 | -7.0943
3.864 | -34.142 | -7.1052
3.8701 | -34.1423 | -7.1162
3.8762 | -34.1425 | -7.1267
3.8823 | -34.1428 | -7.1376
3.8884 | -34.1431 | -7.1483
3.8945 | -34.1435 | -7.1592
3.9006 | -34.1437 | -7.1703
3.9067 | -34.144 | -7.1814
3.9128 | -34.1443 | -7.1922
3.9189 | -34.1446 | -7.2028
3.925 | -34.1448 | -7.2137
3.9311 | -34.1451 | -7.224
3.9372 | -34.1454 | -7.2349
3.9434 | -34.1457 | -7.2456
3.9495 | -34.1461 | -7.2565
3.9556 | -34.1464 | -7.267
3.9617 | -34.1467 | -7.2774
3.9678 | -34.147 | -7.2879
3.9739 | -34.1473 | -7.2986
3.98 | -34.1476 | -7.3094
3.9861 | -34.1479 | -7.3201
3.9922 | -34.1481 | -7.3308
3.9983 | -34.1484 | -7.3412
4.0044 | -34.1487 | -7.3515
4.0105 | -34.149 | -7.3622
4.0166 | -34.1493 | -7.373
4.0227 | -34.1496 | -7.3837
4.0288 | -34.1498 | -7.3948
4.0349 | -34.1501 | -7.4053
4.041 | -34.1503 | -7.416
4.0471 | -34.1505 | -7.4267
4.0532 | -34.1508 | -7.4371
4.0593 | -34.1512 | -7.448
4.0654 | -34.1514 | -7.4585
4.0715 | -34.1517 | -7.4696
4.0776 | -34.152 | -7.4802
4.0838 | -34.1523 | -7.4914
4.0899 | -34.1526 | -7.5018
4.096 | -34.1529 | -7.5124
4.1021 | -34.1531 | -7.5229
4.1082 | -34.1534 | -7.5333
4.1143 | -34.1537 | -7.5437
4.1204 | -34.154 | -7.554
4.1265 | -34.1542 | -7.5644
4.1326 | -34.1545 | -7.5748
4.1387 | -34.1548 | -7.5853
4.1448 | -34.1551 | -7.5961
4.1509 | -34.1553 | -7.6067
4.157 | -34.1555 | -7.6177
4.1631 | -34.1558 | -7.6285
4.1692 | -34.1561 | -7.639
4.1753 | -34.1564 | -7.6496
4.1814 | -34.1566 | -7.6598
4.1875 | -34.1569 | -7.6707
4.1936 | -34.1571 | -7.6814
4.1997 | -34.1573 | -7.6916
4.2058 | -34.1576 | -7.7024
4.2119 | -34.1578 | -7.7129
4.218 | -34.1581 | -7.7231
4.2241 | -34.1583 | -7.7335
4.2303 | -34.1586 | -7.7439
4.2364 | -34.1588 | -7.7545
4.2425 | -34.159 | -7.7651
4.2486 | -34.1592 | -7.7758
4.2547 | -34.1594 | -7.7861
4.2608 | -34.1596 | -7.7967
4.2669 | -34.1598 | -7.8069
4.273 | -34.16 | -7.8176
4.2791 | -34.1603 | -7.8281
4.2852 | -34.1605 | -7.8388
4.2913 | -34.1607 | -7.849
4.2974 | -34.1609 | -7.8598
4.3035 | -34.1612 | -7.8706
4.3096 | -34.1613 | -7.8813
4.3157 | -34.1616 | -7.8916
4.3218 | -34.1618 | -7.902
4.3279 | -34.162 | -7.9124
4.334 | -34.1622 | -7.9228
4.3401 | -34.1624 | -7.9329
4.3462 | -34.1627 | -7.9435
4.3523 | -34.1629 | -7.9536
4.3584 | -34.1631 | -7.9637
4.3645 | -34.1633 | -7.974
4.3707 | -34.1635 | -7.9842
4.3768 | -34.1637 | -7.9946
4.3829 | -34.164 | -8.0053
4.389 | -34.1642 | -8.0157
4.3951 | -34.1644 | -8.0259
4.4012 | -34.1646 | -8.0366
4.4073 | -34.1648 | -8.0468
4.4134 | -34.165 | -8.0571
4.4195 | -34.1653 | -8.0674
4.4256 | -34.1655 | -8.0776
4.4317 | -34.1657 | -8.0879
4.4378 | -34.1659 | -8.098
4.4439 | -34.1661 | -8.108
4.45 | -34.1663 | -8.1183
4.4561 | -34.1664 | -8.1287
4.4622 | -34.1666 | -8.139
4.4683 | -34.1669 | -8.1493
4.4744 | -34.1671 | -8.16
4.4805 | -34.1673 | -8.1699
4.4866 | -34.1675 | -8.1802
4.4927 | -34.1677 | -8.1908
4.4988 | -34.1679 | -8.201
4.5049 | -34.1682 | -8.2109
4.511 | -34.1684 | -8.221
4.5172 | -34.1686 | -8.2313
4.5233 | -34.1688 | -8.2419
4.5294 | -34.1691 | -8.2521
4.5355 | -34.1693 | -8.2624
4.5416 | -34.1695 | -8.2726
4.5477 | -34.1697 | -8.2829
4.5538 | -34.1699 | -8.2928
4.5599 | -34.1702 | -8.3031
4.566 | -34.1704 | -8.3133
4.5721 | -34.1706 | -8.3236
4.5782 | -34.1709 | -8.3341
4.5843 | -34.1711 | -8.3444
4.5904 | -34.1714 | -8.3548
4.5965 | -34.1716 | -8.3649
4.6026 | -34.1718 | -8.3756
4.6087 | -34.1721 | -8.3858
4.6148 | -34.1723 | -8.3964
4.6209 | -34.1725 | -8.4066
4.627 | -34.1728 | -8.417
4.6331 | -34.1729 | -8.4272
4.6392 | -34.1732 | -8.4375
4.6453 | -34.1734 | -8.4478
4.6514 | -34.1737 | -8.4581
4.6576 | -34.1739 | -8.4688
4.6637 | -34.174 | -8.4788
4.6698 | -34.1742 | -8.4889
4.6759 | -34.1745 | -8.4996
4.682 | -34.1747 | -8.5098
4.6881 | -34.1749 | -8.5201
4.6942 | -34.1751 | -8.5303
4.7003 | -34.1753 | -8.5406
4.7064 | -34.1756 | -8.5508
4.7125 | -34.1758 | -8.5614
4.7186 | -34.176 | -8.5715
4.7247 | -34.1763 | -8.5818
4.7308 | -34.1765 | -8.5921
4.7369 | -34.1768 | -8.6025
4.743 | -34.177 | -8.6132
4.7491 | -34.1773 | -8.6238
4.7552 | -34.1775 | -8.6343
4.7613 | -34.1777 | -8.6447
4.7674 | -34.1779 | -8.6549
4.7735 | -34.1781 | -8.6656
4.7796 | -34.1784 | -8.6762
4.7857 | -34.1786 | -8.6867
4.7918 | -34.1789 | -8.6968
4.7979 | -34.1792 | -8.7071
4.8041 | -34.1794 | -8.7175
4.8102 | -34.1797 | -8.7275
4.8163 | -34.1799 | -8.7378
4.8224 | -34.1801 | -8.7481
4.8285 | -34.1804 | -8.7582
4.8346 | -34.1807 | -8.7684
4.8407 | -34.1809 | -8.7788
4.8468 | -34.1812 | -8.7894
4.8529 | -34.1814 | -8.8001
4.859 | -34.1816 | -8.8107
4.8651 | -34.1819 | -8.8209
4.8712 | -34.1821 | -8.8311
4.8773 | -34.1823 | -8.8415
4.8834 | -34.1826 | -8.8519
4.8895 | -34.1829 | -8.8621
4.8956 | -34.1831 | -8.8724
4.9017 | -34.1833 | -8.8827
4.9078 | -34.1835 | -8.8933
4.9139 | -34.1838 | -8.9034
4.92 | -34.1841 | -8.9134
4.9261 | -34.1844 | -8.9235
4.9322 | -34.1846 | -8.9338
4.9383 | -34.1849 | -8.9441
4.9445 | -34.1852 | -8.9545
4.9506 | -34.1854 | -8.9647
4.9567 | -34.1857 | -8.9749
4.9628 | -34.1859 | -8.9852
4.9689 | -34.1862 | -8.996
4.975 | -34.1864 | -9.0066
4.9811 | -34.1867 | -9.0167
4.9872 | -34.187 | -9.0269
4.9933 | -34.1873 | -9.0375
4.9994 | -34.1876 | -9.0482
5.0055 | -34.1879 | -9.0586
5.0116 | -34.1881 | -9.0694
5.0177 | -34.1885 | -9.08
5.0238 | -34.1888 | -9.0907
5.0299 | -34.1891 | -9.1014
5.036 | -34.1894 | -9.1118
5.0421 | -34.1896 | -9.122
5.0482 | -34.1899 | -9.1324
5.0543 | -34.1902 | -9.1426
5.0604 | -34.1905 | -9.153
5.0665 | -34.1908 | -9.1631
5.0726 | -34.1911 | -9.1739
5.0787 | -34.1914 | -9.1844
5.0848 | -34.1917 | -9.195
5.091 | -34.192 | -9.2052
5.0971 | -34.1923 | -9.2153
5.1032 | -34.1926 | -9.226
5.1093 | -34.1929 | -9.2364
5.1154 | -34.1931 | -9.2466
5.1215 | -34.1934 | -9.2569
5.1276 | -34.1937 | -9.2671
5.1337 | -34.194 | -9.2772
5.1398 | -34.1943 | -9.2874
5.1459 | -34.1946 | -9.2981
5.152 | -34.1949 | -9.3086
5.1581 | -34.1951 | -9.3192
5.1642 | -34.1955 | -9.3296
5.1703 | -34.1957 | -9.3399
5.1764 | -34.1961 | -9.3502
5.1825 | -34.1964 | -9.3603
5.1886 | -34.1967 | -9.3704
5.1947 | -34.1969 | -9.3808
5.2008 | -34.1973 | -9.3913
5.2069 | -34.1975 | -9.4018
5.213 | -34.1978 | -9.412
5.2191 | -34.1982 | -9.4225
5.2252 | -34.1984 | -9.4329
5.2314 | -34.1987 | -9.4435
5.2375 | -34.199 | -9.4536
5.2436 | -34.1993 | -9.4641
5.2497 | -34.1996 | -9.4741
5.2558 | -34.1999 | -9.4849
5.2619 | -34.2002 | -9.4955
5.268 | -34.2005 | -9.5061
5.2741 | -34.2009 | -9.5165
5.2802 | -34.2012 | -9.5273
5.2863 | -34.2015 | -9.5376
5.2924 | -34.2018 | -9.5482
5.2985 | -34.2021 | -9.5587
5.3046 | -34.2024 | -9.5691
5.3107 | -34.2027 | -9.5797
5.3168 | -34.203 | -9.5902
5.3229 | -34.2033 | -9.6006
5.329 | -34.2036 | -9.6111
5.3351 | -34.204 | -9.6212
5.3412 | -34.2042 | -9.6316
5.3473 | -34.2046 | -9.642
5.3534 | -34.2049 | -9.6525
5.3595 | -34.2052 | -9.663
5.3656 | -34.2055 | -9.6733
5.3717 | -34.2058 | -9.684
5.3779 | -34.2061 | -9.6948
5.384 | -34.2064 | -9.7052
5.3901 | -34.2067 | -9.7159
5.3962 | -34.207 | -9.7261
5.4023 | -34.2074 | -9.7368
5.4084 | -34.2077 | -9.7471
5.4145 | -34.208 | -9.7573
5.4206 | -34.2083 | -9.7678
5.4267 | -34.2086 | -9.7785
5.4328 | -34.2089 | -9.789
5.4389 | -34.2093 | -9.7991
5.445 | -34.2095 | -9.8099
5.4511 | -34.2099 | -9.8201
5.4572 | -34.2102 | -9.8307
5.4633 | -34.2105 | -9.8411
5.4694 | -34.2108 | -9.8516
5.4755 | -34.2111 | -9.8618
5.4816 | -34.2114 | -9.8723
5.4877 | -34.2117 | -9.8826
5.4938 | -34.212 | -9.8928
5.4999 | -34.2123 | -9.9031
5.506 | -34.2127 | -9.9135
5.5121 | -34.2129 | -9.9238
5.5183 | -34.2133 | -9.9339
5.5244 | -34.2135 | -9.9446
5.5305 | -34.2139 | -9.9549
5.5366 | -34.2142 | -9.9651
5.5427 | -34.2144 | -9.9756
5.5488 | -34.2147 | -9.9861
5.5549 | -34.215 | -9.9965
5.561 | -34.2154 | -10.0071
5.5671 | -34.2157 | -10.0174
5.5732 | -34.216 | -10.028
5.5793 | -34.2163 | -10.0387
5.5854 | -34.2167 | -10.049
5.5915 | -34.2169 | -10.0597
5.5976 | -34.2172 | -10.0702
5.6037 | -34.2175 | -10.0807
5.6098 | -34.2178 | -10.0911
5.6159 | -34.2181 | -10.1013
5.622 | -34.2184 | -10.1118
5.6281 | -34.2188 | -10.1222
5.6342 | -34.2191 | -10.1329
5.6403 | -34.2194 | -10.1434
5.6464 | -34.2197 | -10.154
5.6525 | -34.22 | -10.1645
5.6586 | -34.2203 | -10.1746
5.6648 | -34.2206 | -10.1851
5.6709 | -34.221 | -10.1957
5.677 | -34.2213 | -10.2062
5.6831 | -34.2216 | -10.2167
5.6892 | -34.2218 | -10.2273
5.6953 | -34.2222 | -10.2376
5.7014 | -34.2225 | -10.2481
5.7075 | -34.2227 | -10.2586
5.7136 | -34.223 | -10.2689
5.7197 | -34.2233 | -10.2793
5.7258 | -34.2236 | -10.2895
5.7319 | -34.2239 | -10.3
5.738 | -34.2242 | -10.3104
5.7441 | -34.2246 | -10.3208
5.7502 | -34.2249 | -10.331
5.7563 | -34.2252 | -10.3414
5.7624 | -34.2256 | -10.3521
5.7685 | -34.2259 | -10.3625
5.7746 | -34.2261 | -10.3732
5.7807 | -34.2264 | -10.3836
5.7868 | -34.2267 | -10.3942
5.7929 | -34.2269 | -10.4046
5.799 | -34.2273 | -10.4148
5.8052 | -34.2276 | -10.4251
5.8113 | -34.2279 | -10.4354
5.8174 | -34.2281 | -10.4456
5.8235 | -34.2285 | -10.456
5.8296 | -34.2287 | -10.4665
5.8357 | -34.229 | -10.4768
5.8418 | -34.2293 | -10.4874
5.8479 | -34.2296 | -10.4979
5.854 | -34.2299 | -10.5084
5.8601 | -34.2302 | -10.5187
5.8662 | -34.2305 | -10.5293
5.8723 | -34.2308 | -10.5395
5.8784 | -34.2311 | -10.5498
5.8845 | -34.2314 | -10.5601
5.8906 | -34.2317 | -10.5706
5.8967 | -34.232 | -10.5808
5.9028 | -34.2323 | -10.591
5.9089 | -34.2327 | -10.6013
5.915 | -34.233 | -10.6115
5.9211 | -34.2333 | -10.6218
5.9272 | -34.2336 | -10.6319
5.9333 | -34.2339 | -10.6421
5.9394 | -34.2342 | -10.6525
5.9455 | -34.2345 | -10.6626
5.9517 | -34.2349 | -10.6728
5.9578 | -34.2352 | -10.6833
5.9639 | -34.2355 | -10.6933
5.97 | -34.2359 | -10.7038
5.9761 | -34.2362 | -10.7141
5.9822 | -34.2365 | -10.7246
5.9883 | -34.2368 | -10.735
5.9944 | -34.2371 | -10.7454
6.0005 | -34.2375 | -10.7559
6.0066 | -34.2378 | -10.7661
6.0127 | -34.2381 | -10.7761
6.0188 | -34.2384 | -10.7867
6.0249 | -34.2387 | -10.7971
6.031 | -34.239 | -10.8072
6.0371 | -34.2393 | -10.8176
6.0432 | -34.2396 | -10.8278
6.0493 | -34.24 | -10.838
6.0554 | -34.2403 | -10.8482
6.0615 | -34.2406 | -10.8586
6.0676 | -34.241 | -10.869
6.0737 | -34.2413 | -10.8795
6.0798 | -34.2416 | -10.8898
6.0859 | -34.2419 | -10.9001
6.0921 | -34.2422 | -10.9102
6.0982 | -34.2425 | -10.9205
6.1043 | -34.2429 | -10.9306
6.1104 | -34.2432 | -10.9412
6.1165 | -34.2436 | -10.9511
6.1226 | -34.2439 | -10.9613
6.1287 | -34.2442 | -10.9715
6.1348 | -34.2445 | -10.9818
6.1409 | -34.2449 | -10.9921
6.147 | -34.2452 | -11.0023
6.1531 | -34.2455 | -11.0126
6.1592 | -34.2458 | -11.0227
6.1653 | -34.2461 | -11.0329
6.1714 | -34.2464 | -11.0433
6.1775 | -34.2468 | -11.0536
6.1836 | -34.2471 | -11.0639
6.1897 | -34.2474 | -11.0742
6.1958 | -34.2477 | -11.0848
6.2019 | -34.2481 | -11.0951
6.208 | -34.2484 | -11.1054
6.2141 | -34.2487 | -11.1156
6.2202 | -34.2491 | -11.126
6.2263 | -34.2494 | -11.1362
6.2325 | -34.2497 | -11.1463
6.2386 | -34.2501 | -11.1563
6.2447 | -34.2504 | -11.1668
6.2508 | -34.2507 | -11.177
6.2569 | -34.2511 | -11.1873
6.263 | -34.2514 | -11.1975
6.2691 | -34.2517 | -11.2077
6.2752 | -34.252 | -11.218
6.2813 | -34.2523 | -11.2282
6.2874 | -34.2526 | -11.2384
6.2935 | -34.253 | -11.2486
6.2996 | -34.2533 | -11.2588
6.3057 | -34.2536 | -11.2691
6.3118 | -34.2539 | -11.2795
6.3179 | -34.2542 | -11.2897
6.324 | -34.2545 | -11.3
6.3301 | -34.2549 | -11.3103
6.3362 | -34.2552 | -11.3204
6.3423 | -34.2556 | -11.3309
6.3484 | -34.2559 | -11.3412
6.3545 | -34.2562 | -11.3512
6.3606 | -34.2565 | -11.3612
6.3667 | -34.2568 | -11.3713
6.3728 | -34.2571 | -11.3816
6.379 | -34.2574 | -11.3918
6.3851 | -34.2578 | -11.4018
6.3912 | -34.2581 | -11.4121
6.3973 | -34.2584 | -11.4223
6.4034 | -34.2587 | -11.4324
6.4095 | -34.259 | -11.4426
6.4156 | -34.2593 | -11.4529
6.4217 | -34.2596 | -11.4631
6.4278 | -34.26 | -11.4733
6.4339 | -34.2603 | -11.4835
6.44 | -34.2607 | -11.4936
6.4461 | -34.261 | -11.504
6.4522 | -34.2613 | -11.514
6.4583 | -34.2616 | -11.5243
6.4644 | -34.2619 | -11.5346
6.4705 | -34.2622 | -11.5446
6.4766 | -34.2625 | -11.5548
6.4827 | -34.2628 | -11.565
6.4888 | -34.2631 | -11.5751
6.4949 | -34.2635 | -11.585
6.501 | -34.2638 | -11.5953
6.5071 | -34.2641 | -11.6053
6.5132 | -34.2645 | -11.6156
6.5194 | -34.2648 | -11.6258
6.5255 | -34.2651 | -11.6358
6.5316 | -34.2654 | -11.6461
6.5377 | -34.2657 | -11.6563
6.5438 | -34.266 | -11.6663
6.5499 | -34.2663 | -11.6764
6.556 | -34.2667 | -11.6865
6.5621 | -34.267 | -11.6969
6.5682 | -34.2673 | -11.7072
6.5743 | -34.2677 | -11.7173
6.5804 | -34.268 | -11.7276
6.5865 | -34.2683 | -11.7378
6.5926 | -34.2686 | -11.7479
6.5987 | -34.2689 | -11.7578
6.6048 | -34.2693 | -11.7681
6.6109 | -34.2696 | -11.7783
6.617 | -34.2699 | -11.7886
6.6231 | -34.2702 | -11.7988
6.6292 | -34.2705 | -11.809
6.6353 | -34.2708 | -11.8195
6.6414 | -34.2712 | -11.8295
6.6475 | -34.2715 | -11.8398
6.6536 | -34.2718 | -11.8499
6.6597 | -34.2721 | -11.8598
6.6659 | -34.2724 | -11.87
6.672 | -34.2728 | -11.8801
6.6781 | -34.2731 | -11.8903
6.6842 | -34.2734 | -11.9003
6.6903 | -34.2738 | -11.9104
6.6964 | -34.2741 | -11.9206
6.7025 | -34.2744 | -11.9305
6.7086 | -34.2747 | -11.9407
6.7147 | -34.275 | -11.9509
6.7208 | -34.2754 | -11.9611
6.7269 | -34.2757 | -11.9712
6.733 | -34.276 | -11.9812
6.7391 | -34.2763 | -11.9913
6.7452 | -34.2767 | -12.0013
6.7513 | -34.277 | -12.0112
6.7574 | -34.2773 | -12.0214
6.7635 | -34.2776 | -12.0313
6.7696 | -34.278 | -12.0413
6.7757 | -34.2783 | -12.0515
6.7818 | -34.2786 | -12.0616
6.7879 | -34.279 | -12.0716
6.794 | -34.2793 | -12.0816
6.8001 | -34.2796 | -12.0917
6.8063 | -34.2799 | -12.1018
6.8124 | -34.2802 | -12.1121
6.8185 | -34.2805 | -12.1224
6.8246 | -34.2808 | -12.1324
6.8307 | -34.2811 | -12.1424
6.8368 | -34.2815 | -12.1527
6.8429 | -34.2818 | -12.1629
6.849 | -34.2821 | -12.173
6.8551 | -34.2825 | -12.183
6.8612 | -34.2828 | -12.1931
6.8673 | -34.2831 | -12.2032
6.8734 | -34.2834 | -12.2135
6.8795 | -34.2838 | -12.2236
6.8856 | -34.2841 | -12.2337
6.8917 | -34.2844 | -12.2437
6.8978 | -34.2847 | -12.2536
6.9039 | -34.285 | -12.2634
6.91 | -34.2854 | -12.2736
6.9161 | -34.2857 | -12.2835
6.9222 | -34.286 | -12.2936
6.9283 | -34.2864 | -12.3034
6.9344 | -34.2867 | -12.3134
6.9405 | -34.2871 | -12.3234
6.9466 | -34.2874 | -12.3337
6.9528 | -34.2878 | -12.3439
6.9589 | -34.2881 | -12.3541
6.965 | -34.2884 | -12.3643
6.9711 | -34.2888 | -12.3745
6.9772 | -34.2892 | -12.3845
6.9833 | -34.2895 | -12.3946
6.9894 | -34.2898 | -12.4048
6.9955 | -34.2901 | -12.4147
7.0016 | -34.2905 | -12.4247
7.0077 | -34.2908 | -12.4347
7.0138 | -34.2911 | -12.4447
7.0199 | -34.2915 | -12.4547
7.026 | -34.2918 | -12.4647
7.0321 | -34.2921 | -12.4746
7.0382 | -34.2925 | -12.4844
7.0443 | -34.2928 | -12.4945
7.0504 | -34.2931 | -12.5046
7.0565 | -34.2935 | -12.5147
7.0626 | -34.2938 | -12.5245
7.0687 | -34.2941 | -12.5342
7.0748 | -34.2945 | -12.5439
7.0809 | -34.2948 | -12.5539
7.087 | -34.2952 | -12.564
7.0932 | -34.2955 | -12.5741
7.0993 | -34.2958 | -12.584
7.1054 | -34.2961 | -12.5939
7.1115 | -34.2965 | -12.6037
7.1176 | -34.2968 | -12.6137
7.1237 | -34.2971 | -12.6236
7.1298 | -34.2975 | -12.6335
7.1359 | -34.2978 | -12.6435
7.142 | -34.2981 | -12.6533
7.1481 | -34.2984 | -12.6631
7.1542 | -34.2988 | -12.6731
7.1603 | -34.2991 | -12.6831
7.1664 | -34.2995 | -12.693
7.1725 | -34.2998 | -12.703
7.1786 | -34.3002 | -12.7129
7.1847 | -34.3005 | -12.7229
7.1908 | -34.3008 | -12.7331
7.1969 | -34.3011 | -12.7432
7.203 | -34.3015 | -12.7529
7.2091 | -34.3018 | -12.7628
7.2152 | -34.3021 | -12.7727
7.2213 | -34.3025 | -12.7825
7.2274 | -34.3028 | -12.7924
7.2335 | -34.3031 | -12.8023
7.2397 | -34.3035 | -12.8122
7.2458 | -34.3039 | -12.8218
7.2519 | -34.3042 | -12.8318
7.258 | -34.3045 | -12.8419
7.2641 | -34.3049 | -12.8518
7.2702 | -34.3052 | -12.8619
7.2763 | -34.3055 | -12.8718
7.2824 | -34.3059 | -12.8818
7.2885 | -34.3062 | -12.8917
7.2946 | -34.3066 | -12.9017
7.3007 | -34.3069 | -12.9116
7.3068 | -34.3072 | -12.9217
7.3129 | -34.3075 | -12.9317
7.319 | -34.3079 | -12.9416
7.3251 | -34.3082 | -12.9514
7.3312 | -34.3086 | -12.9613
7.3373 | -34.3089 | -12.971
7.3434 | -34.3093 | -12.9809
7.3495 | -34.3096 | -12.9906
7.3556 | -34.31 | -13.0006
7.3617 | -34.3103 | -13.0107
7.3678 | -34.3107 | -13.0205
7.3739 | -34.311 | -13.0306
7.3801 | -34.3113 | -13.0407
7.3862 | -34.3117 | -13.0505
7.3923 | -34.312 | -13.0606
7.3984 | -34.3124 | -13.0703
7.4045 | -34.3127 | -13.0801
7.4106 | -34.3131 | -13.0899
7.4167 | -34.3134 | -13.1
7.4228 | -34.3138 | -13.1099
7.4289 | -34.3141 | -13.1198
7.435 | -34.3145 | -13.1294
7.4411 | -34.3148 | -13.1393
7.4472 | -34.3151 | -13.1493
7.4533 | -34.3155 | -13.159
7.4594 | -34.3159 | -13.1691
7.4655 | -34.3162 | -13.1788
7.4716 | -34.3166 | -13.1889
7.4777 | -34.3169 | -13.1987
7.4838 | -34.3172 | -13.2084
7.4899 | -34.3176 | -13.2181
7.496 | -34.3179 | -13.2278
7.5021 | -34.3183 | -13.2377
7.5082 | -34.3186 | -13.2476
7.5143 | -34.3189 | -13.2575
7.5204 | -34.3193 | -13.2675
7.5266 | -34.3196 | -13.2772
7.5327 | -34.32 | -13.287
7.5388 | -34.3203 | -13.2971
7.5449 | -34.3207 | -13.3072
7.551 | -34.321 | -13.3169
7.5571 | -34.3214 | -13.3269
7.5632 | -34.3217 | -13.3368
7.5693 | -34.3221 | -13.3465
7.5754 | -34.3224 | -13.3562
7.5815 | -34.3228 | -13.3661
7.5876 | -34.3231 | -13.3762
7.5937 | -34.3235 | -13.3861
7.5998 | -34.3239 | -13.3959
7.6059 | -34.3242 | -13.406
7.612 | -34.3246 | -13.4159
7.6181 | -34.3249 | -13.4257
7.6242 | -34.3253 | -13.4358
7.6303 | -34.3256 | -13.4458
7.6364 | -34.3259 | -13.4558
7.6425 | -34.3263 | -13.4657
7.6486 | -34.3266 | -13.4756
7.6547 | -34.327 | -13.4854
7.6608 | -34.3273 | -13.4952
7.667 | -34.3277 | -13.5051
7.6731 | -34.328 | -13.5151
7.6792 | -34.3284 | -13.5252
7.6853 | -34.3287 | -13.5352
7.6914 | -34.3291 | -13.5451
7.6975 | -34.3294 | -13.5551
7.7036 | -34.3298 | -13.565
7.7097 | -34.3302 | -13.575
7.7158 | -34.3305 | -13.5849
7.7219 | -34.3309 | -13.5947
7.728 | -34.3312 | -13.6046
7.7341 | -34.3316 | -13.6146
7.7402 | -34.3319 | -13.6245
7.7463 | -34.3323 | -13.6344
7.7524 | -34.3326 | -13.6446
7.7585 | -34.333 | -13.6545
7.7646 | -34.3333 | -13.6644
7.7707 | -34.3337 | -13.6744
7.7768 | -34.334 | -13.6845
7.7829 | -34.3344 | -13.6946
7.789 | -34.3347 | -13.7047
7.7951 | -34.3351 | -13.7146
7.8012 | -34.3354 | -13.7245
7.8073 | -34.3358 | -13.7346
7.8135 | -34.3362 | -13.7445
7.8196 | -34.3365 | -13.7546
7.8257 | -34.3369 | -13.7647
7.8318 | -34.3372 | -13.7747
7.8379 | -34.3376 | -13.7845
7.844 | -34.3379 | -13.7944
7.8501 | -34.3383 | -13.8045
7.8562 | -34.3386 | -13.8147
7.8623 | -34.339 | -13.8245
7.8684 | -34.3393 | -13.8345
7.8745 | -34.3397 | -13.8445
7.8806 | -34.34 | -13.8546
7.8867 | -34.3404 | -13.8644
7.8928 | -34.3408 | -13.8741
7.8989 | -34.3411 | -13.8841
7.905 | -34.3415 | -13.8941
7.9111 | -34.3418 | -13.9041
7.9172 | -34.3421 | -13.9141
7.9233 | -34.3425 | -13.9241
7.9294 | -34.3428 | -13.9342
7.9355 | -34.3432 | -13.9439
7.9416 | -34.3435 | -13.9539
7.9477 | -34.3439 | -13.9638
7.9539 | -34.3442 | -13.9736
7.96 | -34.3446 | -13.9836
7.9661 | -34.3449 | -13.9936
7.9722 | -34.3453 | -14.0034
7.9783 | -34.3456 | -14.0136
7.9844 | -34.346 | -14.0235
7.9905 | -34.3464 | -14.0336
7.9966 | -34.3467 | -14.0437
8.0027 | -34.3471 | -14.0536
8.0088 | -34.3474 | -14.0636
8.0149 | -34.3477 | -14.0735
8.021 | -34.3481 | -14.0836
8.0271 | -34.3485 | -14.0938
8.0332 | -34.3488 | -14.1036
8.0393 | -34.3492 | -14.1137
8.0454 | -34.3495 | -14.1238
8.0515 | -34.3499 | -14.1338
8.0576 | -34.3503 | -14.144
8.0637 | -34.3506 | -14.1538
8.0698 | -34.351 | -14.1637
8.0759 | -34.3513 | -14.1736
8.082 | -34.3517 | -14.1835
8.0881 | -34.352 | -14.1935
8.0942 | -34.3524 | -14.2035
8.1004 | -34.3528 | -14.2135
8.1065 | -34.3531 | -14.2234
8.1126 | -34.3534 | -14.2334
8.1187 | -34.3538 | -14.2435
8.1248 | -34.3541 | -14.2533
8.1309 | -34.3545 | -14.2631
8.137 | -34.3549 | -14.2729
8.1431 | -34.3552 | -14.2831
8.1492 | -34.3556 | -14.293
8.1553 | -34.3559 | -14.303
8.1614 | -34.3563 | -14.3132
8.1675 | -34.3566 | -14.3231
8.1736 | -34.357 | -14.3332
8.1797 | -34.3573 | -14.3433
8.1858 | -34.3577 | -14.3533
8.1919 | -34.358 | -14.3633
8.198 | -34.3584 | -14.3731
8.2041 | -34.3588 | -14.3832
8.2102 | -34.3591 | -14.3931
8.2163 | -34.3595 | -14.403
8.2224 | -34.3598 | -14.413
8.2285 | -34.3602 | -14.4227
8.2346 | -34.3605 | -14.4327
8.2408 | -34.3609 | -14.4425
8.2469 | -34.3613 | -14.4526
8.253 | -34.3616 | -14.4625
8.2591 | -34.362 | -14.4726
8.2652 | -34.3624 | -14.4826
8.2713 | -34.3627 | -14.4927
8.2774 | -34.3631 | -14.5027
8.2835 | -34.3634 | -14.5129
8.2896 | -34.3638 | -14.5227
8.2957 | -34.3641 | -14.5328
8.3018 | -34.3645 | -14.5428
8.3079 | -34.3649 | -14.5527
8.314 | -34.3653 | -14.5626
8.3201 | -34.3656 | -14.5725
8.3262 | -34.366 | -14.5826
8.3323 | -34.3663 | -14.5925
8.3384 | -34.3667 | -14.6026
8.3445 | -34.367 | -14.6125
8.3506 | -34.3674 | -14.6225
8.3567 | -34.3678 | -14.6325
8.3628 | -34.3681 | -14.6425
8.3689 | -34.3685 | -14.6524
8.375 | -34.3688 | -14.6624
8.3812 | -34.3692 | -14.6722
8.3873 | -34.3696 | -14.6822
8.3934 | -34.3699 | -14.6923
8.3995 | -34.3702 | -14.7021
8.4056 | -34.3706 | -14.712
8.4117 | -34.371 | -14.7221
8.4178 | -34.3713 | -14.732
8.4239 | -34.3717 | -14.7418
8.43 | -34.372 | -14.7518
8.4361 | -34.3723 | -14.7618
8.4422 | -34.3727 | -14.7717
8.4483 | -34.3731 | -14.7816
8.4544 | -34.3735 | -14.7916
8.4605 | -34.3738 | -14.8014
8.4666 | -34.3742 | -14.8115
8.4727 | -34.3745 | -14.8215
8.4788 | -34.3749 | -14.8315
8.4849 | -34.3752 | -14.8414
8.491 | -34.3756 | -14.8514
8.4971 | -34.376 | -14.8611
8.5032 | -34.3763 | -14.871
8.5093 | -34.3767 | -14.8808
8.5154 | -34.3771 | -14.8908
8.5215 | -34.3775 | -14.9006
8.5277 | -34.3778 | -14.9106
8.5338 | -34.3782 | -14.9205
8.5399 | -34.3785 | -14.9305
8.546 | -34.3789 | -14.9403
8.5521 | -34.3793 | -14.9502
8.5582 | -34.3796 | -14.9602
8.5643 | -34.3799 | -14.9699
8.5704 | -34.3803 | -14.9795
8.5765 | -34.3807 | -14.9896
8.5826 | -34.381 | -14.9992
8.5887 | -34.3814 | -15.009
8.5948 | -34.3818 | -15.019
8.6009 | -34.3822 | -15.0287
8.607 | -34.3825 | -15.0387
8.6131 | -34.3829 | -15.0486
8.6192 | -34.3833 | -15.0583
8.6253 | -34.3837 | -15.0681
8.6314 | -34.384 | -15.078
8.6375 | -34.3844 | -15.088
8.6436 | -34.3848 | -15.0979
8.6497 | -34.3851 | -15.1077
8.6558 | -34.3855 | -15.1176
8.6619 | -34.3858 | -15.1275
8.6681 | -34.3862 | -15.1375
8.6742 | -34.3866 | -15.1474
8.6803 | -34.3869 | -15.1572
8.6864 | -34.3873 | -15.1668
8.6925 | -34.3877 | -15.1765
8.6986 | -34.388 | -15.1865
8.7047 | -34.3884 | -15.1965
8.7108 | -34.3888 | -15.2064
8.7169 | -34.3891 | -15.2163
8.723 | -34.3895 | -15.2261
8.7291 | -34.3899 | -15.2358
8.7352 | -34.3902 | -15.2458
8.7413 | -34.3906 | -15.2558
8.7474 | -34.3909 | -15.2655
8.7535 | -34.3913 | -15.2754
8.7596 | -34.3917 | -15.2853
8.7657 | -34.3921 | -15.2953
8.7718 | -34.3924 | -15.3053
8.7779 | -34.3928 | -15.3152
8.784 | -34.3932 | -15.3251
8.7901 | -34.3936 | -15.3349
8.7962 | -34.394 | -15.3446
8.8023 | -34.3943 | -15.3544
8.8084 | -34.3947 | -15.3642
8.8146 | -34.3951 | -15.374
8.8207 | -34.3954 | -15.3838
8.8268 | -34.3958 | -15.3934
8.8329 | -34.3962 | -15.4033
8.839 | -34.3966 | -15.4131
8.8451 | -34.3969 | -15.4229
8.8512 | -34.3973 | -15.4326
8.8573 | -34.3976 | -15.4425
8.8634 | -34.398 | -15.4521
8.8695 | -34.3984 | -15.4621
8.8756 | -34.3988 | -15.472
8.8817 | -34.3991 | -15.4818
8.8878 | -34.3995 | -15.4917
8.8939 | -34.3999 | -15.5016
8.9 | -34.4003 | -15.5114
8.9061 | -34.4006 | -15.5212
8.9122 | -34.401 | -15.531
8.9183 | -34.4014 | -15.5409
8.9244 | -34.4018 | -15.5509
8.9305 | -34.4022 | -15.5604
8.9366 | -34.4026 | -15.5702
8.9427 | -34.403 | -15.5801
8.9488 | -34.4034 | -15.5899
8.955 | -34.4037 | -15.5997
8.9611 | -34.4041 | -15.6096
8.9672 | -34.4045 | -15.6194
8.9733 | -34.4049 | -15.629
8.9794 | -34.4053 | -15.6387
8.9855 | -34.4057 | -15.6486
8.9916 | -34.4061 | -15.6586
8.9977 | -34.4065 | -15.6683
9.0038 | -34.4069 | -15.678
9.0099 | -34.4072 | -15.6878
9.016 | -34.4076 | -15.6976
9.0221 | -34.408 | -15.7072
9.0282 | -34.4084 | -15.7168
9.0343 | -34.4088 | -15.7266
9.0404 | -34.4092 | -15.7363
9.0465 | -34.4096 | -15.7462
9.0526 | -34.41 | -15.7559
9.0587 | -34.4104 | -15.7657
9.0648 | -34.4108 | -15.7753
9.0709 | -34.4112 | -15.7851
9.077 | -34.4116 | -15.7949
9.0831 | -34.412 | -15.8048
9.0892 | -34.4123 | -15.8146
9.0953 | -34.4128 | -15.8245
9.1015 | -34.4131 | -15.8343
9.1076 | -34.4135 | -15.8439
9.1137 | -34.4139 | -15.8536
9.1198 | -34.4143 | -15.8633
9.1259 | -34.4147 | -15.8731
9.132 | -34.4151 | -15.883
9.1381 | -34.4155 | -15.8926
9.1442 | -34.416 | -15.902
9.1503 | -34.4163 | -15.9119
9.1564 | -34.4167 | -15.9216
9.1625 | -34.4171 | -15.9314
9.1686 | -34.4175 | -15.941
9.1747 | -34.4179 | -15.9508
9.1808 | -34.4184 | -15.9605
9.1869 | -34.4188 | -15.9702
9.193 | -34.4191 | -15.98
9.1991 | -34.4195 | -15.9898
9.2052 | -34.4199 | -15.9995
9.2113 | -34.4203 | -16.0093
9.2174 | -34.4208 | -16.019
9.2235 | -34.4212 | -16.0287
9.2296 | -34.4216 | -16.0385
9.2357 | -34.422 | -16.0483
9.2419 | -34.4224 | -16.058
9.248 | -34.4228 | -16.0677
9.2541 | -34.4232 | -16.0776
9.2602 | -34.4236 | -16.0874
9.2663 | -34.424 | -16.0969
9.2724 | -34.4244 | -16.1066
9.2785 | -34.4248 | -16.1163
9.2846 | -34.4253 | -16.1259
9.2907 | -34.4257 | -16.1356
9.2968 | -34.4261 | -16.1455
9.3029 | -34.4265 | -16.1553
9.309 | -34.4269 | -16.165
9.3151 | -34.4273 | -16.1746
9.3212 | -34.4277 | -16.1842
9.3273 | -34.4281 | -16.1938
9.3334 | -34.4285 | -16.2034
9.3395 | -34.4289 | -16.213
9.3456 | -34.4293 | -16.2225
9.3517 | -34.4297 | -16.2323
9.3578 | -34.4301 | -16.242
9.3639 | -34.4305 | -16.2515
9.37 | -34.431 | -16.2612
9.3761 | -34.4314 | -16.2711
9.3822 | -34.4318 | -16.2809
9.3884 | -34.4322 | -16.2903
9.3945 | -34.4326 | -16.3001
9.4006 | -34.4331 | -16.3098
9.4067 | -34.4335 | -16.3195
9.4128 | -34.4339 | -16.3292
9.4189 | -34.4343 | -16.339
9.425 | -34.4347 | -16.3486
9.4311 | -34.4351 | -16.3582
9.4372 | -34.4356 | -16.3679
9.4433 | -34.436 | -16.3775
9.4494 | -34.4364 | -16.3871
9.4555 | -34.4368 | -16.3968
9.4616 | -34.4372 | -16.4064
9.4677 | -34.4377 | -16.4161
9.4738 | -34.4381 | -16.4258
9.4799 | -34.4385 | -16.4354
9.486 | -34.4389 | -16.4451
9.4921 | -34.4393 | -16.4547
9.4982 | -34.4397 | -16.4642
9.5043 | -34.4402 | -16.4738
9.5104 | -34.4406 | -16.4834
9.5165 | -34.441 | -16.493
9.5226 | -34.4414 | -16.5028
9.5288 | -34.4418 | -16.5123
9.5349 | -34.4422 | -16.522
9.541 | -34.4426 | -16.5314
9.5471 | -34.443 | -16.541
9.5532 | -34.4434 | -16.5508
9.5593 | -34.4439 | -16.5606
9.5654 | -34.4443 | -16.5701
9.5715 | -34.4447 | -16.5797
9.5776 | -34.4451 | -16.5895
9.5837 | -34.4455 | -16.599
9.5898 | -34.4459 | -16.6088
9.5959 | -34.4463 | -16.6185
9.602 | -34.4467 | -16.6279
9.6081 | -34.4471 | -16.6375
9.6142 | -34.4476 | -16.6471
9.6203 | -34.448 | -16.6567
9.6264 | -34.4484 | -16.6662
9.6325 | -34.4488 | -16.676
9.6386 | -34.4492 | -16.6855
9.6447 | -34.4496 | -16.6952
9.6508 | -34.45 | -16.7047
9.6569 | -34.4504 | -16.7143
9.663 | -34.4509 | -16.724
9.6691 | -34.4513 | -16.7336
9.6753 | -34.4517 | -16.7431
9.6814 | -34.4521 | -16.7526
9.6875 | -34.4525 | -16.762
9.6936 | -34.453 | -16.7715
9.6997 | -34.4534 | -16.7812
9.7058 | -34.4538 | -16.7907
9.7119 | -34.4543 | -16.8004
9.718 | -34.4547 | -16.8101
9.7241 | -34.4551 | -16.8196
9.7302 | -34.4555 | -16.8292
9.7363 | -34.4559 | -16.8387
9.7424 | -34.4564 | -16.8482
9.7485 | -34.4567 | -16.8578
9.7546 | -34.4571 | -16.8672
9.7607 | -34.4576 | -16.877
9.7668 | -34.458 | -16.8866
9.7729 | -34.4584 | -16.8961
9.779 | -34.4589 | -16.9054
9.7851 | -34.4593 | -16.9151
9.7912 | -34.4597 | -16.9246
9.7973 | -34.4602 | -16.9342
9.8034 | -34.4606 | -16.9436
9.8095 | -34.461 | -16.9533
9.8157 | -34.4614 | -16.9628
9.8218 | -34.4618 | -16.9726
9.8279 | -34.4622 | -16.982
9.834 | -34.4626 | -16.9917
9.8401 | -34.4631 | -17.0012
9.8462 | -34.4635 | -17.0105
9.8523 | -34.4639 | -17.0201
9.8584 | -34.4643 | -17.0296
9.8645 | -34.4647 | -17.0393
9.8706 | -34.4651 | -17.049
9.8767 | -34.4655 | -17.0585
9.8828 | -34.4659 | -17.0679
9.8889 | -34.4664 | -17.0772
9.895 | -34.4668 | -17.0868
9.9011 | -34.4672 | -17.0963
9.9072 | -34.4676 | -17.1058
9.9133 | -34.468 | -17.1153
9.9194 | -34.4684 | -17.1248
9.9255 | -34.4688 | -17.1342
9.9316 | -34.4693 | -17.1438
9.9377 | -34.4697 | -17.1532
9.9438 | -34.4701 | -17.1628
9.9499 | -34.4705 | -17.1725
9.956 | -34.4709 | -17.182
9.9622 | -34.4713 | -17.1916
9.9683 | -34.4717 | -17.2012
9.9744 | -34.4721 | -17.2109
9.9805 | -34.4726 | -17.2203
9.9866 | -34.473 | -17.2298
9.9927 | -34.4734 | -17.2392
9.9988 | -34.4738 | -17.2488
10.0049 | -34.4742 | -17.2583
10.011 | -34.4746 | -17.2679
10.0171 | -34.4751 | -17.2773
10.0232 | -34.4755 | -17.2867
10.0293 | -34.4759 | -17.2963
10.0354 | -34.4763 | -17.3058
10.0415 | -34.4768 | -17.3153
10.0476 | -34.4772 | -17.3249
10.0537 | -34.4776 | -17.3345
10.0598 | -34.478 | -17.3439
10.0659 | -34.4784 | -17.3535
10.072 | -34.4788 | -17.3629
10.0781 | -34.4792 | -17.3723
10.0842 | -34.4797 | -17.3818
10.0903 | -34.4801 | -17.3913
10.0964 | -34.4805 | -17.4005
10.1026 | -34.4809 | -17.4101
10.1087 | -34.4813 | -17.4195
10.1148 | -34.4817 | -17.4291
10.1209 | -34.4821 | -17.4385
10.127 | -34.4825 | -17.4478
10.1331 | -34.4829 | -17.4571
10.1392 | -34.4833 | -17.4665
10.1453 | -34.4838 | -17.4759
10.1514 | -34.4842 | -17.4854
10.1575 | -34.4846 | -17.495
10.1636 | -34.485 | -17.5042
10.1697 | -34.4854 | -17.5138
10.1758 | -34.4858 | -17.5231
10.1819 | -34.4862 | -17.5325
10.188 | -34.4866 | -17.542
10.1941 | -34.487 | -17.5514
10.2002 | -34.4874 | -17.5608
10.2063 | -34.4879 | -17.5703
10.2124 | -34.4883 | -17.5798
10.2185 | -34.4887 | -17.5893
10.2246 | -34.4891 | -17.5986
10.2307 | -34.4895 | -17.6081
10.2368 | -34.4899 | -17.6173
10.2429 | -34.4903 | -17.6267
10.2491 | -34.4907 | -17.6362
10.2552 | -34.4911 | -17.6457
10.2613 | -34.4916 | -17.6553
10.2674 | -34.492 | -17.6647
10.2735 | -34.4924 | -17.6744
10.2796 | -34.4928 | -17.6838
10.2857 | -34.4932 | -17.6934
10.2918 | -34.4936 | -17.703
10.2979 | -34.494 | -17.7124
10.304 | -34.4945 | -17.7216
10.3101 | -34.4949 | -17.7312
10.3162 | -34.4953 | -17.7407
10.3223 | -34.4957 | -17.75
10.3284 | -34.4961 | -17.7594
10.3345 | -34.4965 | -17.7688
10.3406 | -34.4969 | -17.7782
10.3467 | -34.4973 | -17.7876
10.3528 | -34.4977 | -17.7969
10.3589 | -34.4981 | -17.8064
10.365 | -34.4985 | -17.8159
10.3711 | -34.4989 | -17.8254
10.3772 | -34.4994 | -17.8348
10.3833 | -34.4998 | -17.8443
10.3895 | -34.5002 | -17.8537
10.3956 | -34.5006 | -17.8631
10.4017 | -34.501 | -17.8724
10.4078 | -34.5014 | -17.8818
10.4139 | -34.5018 | -17.8912
10.42 | -34.5022 | -17.9007
10.4261 | -34.5027 | -17.9102
10.4322 | -34.5031 | -17.9196
10.4383 | -34.5034 | -17.9289
10.4444 | -34.5038 | -17.9382
10.4505 | -34.5042 | -17.9477
10.4566 | -34.5046 | -17.957
10.4627 | -34.505 | -17.9663
10.4688 | -34.5055 | -17.9756
10.4749 | -34.5059 | -17.985
10.481 | -34.5063 | -17.9944
10.4871 | -34.5067 | -18.0037
10.4932 | -34.5071 | -18.013
10.4993 | -34.5075 | -18.0225
10.5054 | -34.5079 | -18.0319
10.5115 | -34.5083 | -18.0413
10.5176 | -34.5088 | -18.0506
10.5237 | -34.5092 | -18.06
10.5298 | -34.5096 | -18.0694
10.536 | -34.51 | -18.0788
10.5421 | -34.5104 | -18.0881
10.5482 | -34.5108 | -18.0975
10.5543 | -34.5113 | -18.1068
10.5604 | -34.5117 | -18.116
10.5665 | -34.5121 | -18.1253
10.5726 | -34.5125 | -18.1346
10.5787 | -34.5129 | -18.144
10.5848 | -34.5133 | -18.1531
10.5909 | -34.5137 | -18.1626
10.597 | -34.5141 | -18.172
10.6031 | -34.5145 | -18.1815
10.6092 | -34.5149 | -18.1909
10.6153 | -34.5153 | -18.2003
10.6214 | -34.5158 | -18.2095
10.6275 | -34.5161 | -18.2189
10.6336 | -34.5166 | -18.2283
10.6397 | -34.5169 | -18.2377
10.6458 | -34.5174 | -18.247
10.6519 | -34.5178 | -18.2565
10.658 | -34.5182 | -18.2658
10.6641 | -34.5186 | -18.2753
10.6702 | -34.519 | -18.2847
10.6764 | -34.5194 | -18.2941
10.6825 | -34.5198 | -18.3033
10.6886 | -34.5202 | -18.3125
10.6947 | -34.5206 | -18.3218
10.7008 | -34.5211 | -18.3312
10.7069 | -34.5215 | -18.3406
10.713 | -34.5219 | -18.35
10.7191 | -34.5223 | -18.3593
10.7252 | -34.5227 | -18.3684
10.7313 | -34.5231 | -18.3776
10.7374 | -34.5235 | -18.387
10.7435 | -34.5239 | -18.3964
10.7496 | -34.5243 | -18.4056
10.7557 | -34.5247 | -18.415
10.7618 | -34.5252 | -18.4242
10.7679 | -34.5256 | -18.4335
10.774 | -34.526 | -18.443
10.7801 | -34.5264 | -18.4523
10.7862 | -34.5268 | -18.4615
10.7923 | -34.5272 | -18.4709
10.7984 | -34.5276 | -18.4801
10.8045 | -34.528 | -18.4892
10.8106 | -34.5284 | -18.4987
10.8168 | -34.5288 | -18.5078
10.8229 | -34.5292 | -18.5171
10.829 | -34.5296 | -18.5264
10.8351 | -34.5301 | -18.5355
10.8412 | -34.5305 | -18.5447
10.8473 | -34.5309 | -18.5538
10.8534 | -34.5313 | -18.5631
10.8595 | -34.5317 | -18.5724
10.8656 | -34.5322 | -18.5815
10.8717 | -34.5326 | -18.5908
10.8778 | -34.533 | -18.6001
10.8839 | -34.5334 | -18.6094
10.89 | -34.5338 | -18.6186
10.8961 | -34.5342 | -18.6279
10.9022 | -34.5346 | -18.6371
10.9083 | -34.5351 | -18.6463
10.9144 | -34.5355 | -18.6555
10.9205 | -34.5359 | -18.6647
10.9266 | -34.5363 | -18.674
10.9327 | -34.5367 | -18.6832
10.9388 | -34.5371 | -18.6923
10.9449 | -34.5375 | -18.7016
10.951 | -34.538 | -18.7109
10.9571 | -34.5384 | -18.7201
10.9633 | -34.5388 | -18.7295
10.9694 | -34.5392 | -18.7389
10.9755 | -34.5396 | -18.7482
10.9816 | -34.54 | -18.7575
10.9877 | -34.5405 | -18.7667
10.9938 | -34.5409 | -18.7761
10.9999 | -34.5413 | -18.7852
11.006 | -34.5417 | -18.7942
11.0121 | -34.5421 | -18.8035
11.0182 | -34.5426 | -18.8128
11.0243 | -34.543 | -18.822
11.0304 | -34.5434 | -18.8313
11.0365 | -34.5438 | -18.8404
11.0426 | -34.5442 | -18.8496
11.0487 | -34.5447 | -18.8589
11.0548 | -34.5451 | -18.868
11.0609 | -34.5455 | -18.8772
11.067 | -34.5459 | -18.8865
11.0731 | -34.5464 | -18.8956
11.0792 | -34.5468 | -18.9048
11.0853 | -34.5472 | -18.914
11.0914 | -34.5476 | -18.9233
11.0975 | -34.5481 | -18.9325
11.1037 | -34.5485 | -18.9417
11.1098 | -34.5489 | -18.9511
11.1159 | -34.5493 | -18.96
11.122 | -34.5497 | -18.9692
11.1281 | -34.5502 | -18.9785
11.1342 | -34.5506 | -18.9877
11.1403 | -34.551 | -18.997
11.1464 | -34.5514 | -19.0063
11.1525 | -34.5519 | -19.0152
11.1586 | -34.5523 | -19.0246
11.1647 | -34.5527 | -19.0337
11.1708 | -34.5531 | -19.0429
11.1769 | -34.5535 | -19.0521
11.183 | -34.554 | -19.0612
11.1891 | -34.5544 | -19.0704
11.1952 | -34.5548 | -19.0796
11.2013 | -34.5553 | -19.0888
11.2074 | -34.5557 | -19.0979
11.2135 | -34.5561 | -19.1073
11.2196 | -34.5565 | -19.1163
11.2257 | -34.557 | -19.1256
11.2318 | -34.5574 | -19.1347
11.2379 | -34.5579 | -19.1439
11.244 | -34.5583 | -19.1533
11.2502 | -34.5587 | -19.1624
11.2563 | -34.5592 | -19.1716
11.2624 | -34.5596 | -19.1806
11.2685 | -34.56 | -19.1898
11.2746 | -34.5604 | -19.1988
11.2807 | -34.5609 | -19.2081
11.2868 | -34.5613 | -19.2174
11.2929 | -34.5618 | -19.2263
11.299 | -34.5622 | -19.2356
11.3051 | -34.5626 | -19.2449
11.3112 | -34.563 | -19.2542
11.3173 | -34.5635 | -19.2632
11.3234 | -34.5639 | -19.2721
11.3295 | -34.5644 | -19.2814
11.3356 | -34.5648 | -19.2907
11.3417 | -34.5652 | -19.2998
11.3478 | -34.5656 | -19.3091
11.3539 | -34.566 | -19.3183
11.36 | -34.5665 | -19.3274
11.3661 | -34.5669 | -19.3364
11.3722 | -34.5673 | -19.3454
11.3783 | -34.5678 | -19.3547
11.3844 | -34.5682 | -19.3639
11.3906 | -34.5686 | -19.3731
11.3967 | -34.5691 | -19.382
11.4028 | -34.5695 | -19.3913
11.4089 | -34.5699 | -19.4003
11.415 | -34.5703 | -19.4095
11.4211 | -34.5708 | -19.4188
11.4272 | -34.5712 | -19.428
11.4333 | -34.5716 | -19.4369
11.4394 | -34.572 | -19.4461
11.4455 | -34.5724 | -19.4551
11.4516 | -34.5729 | -19.4643
11.4577 | -34.5733 | -19.4735
11.4638 | -34.5738 | -19.4823
11.4699 | -34.5742 | -19.4915
11.476 | -34.5746 | -19.5008
11.4821 | -34.575 | -19.5099
11.4882 | -34.5755 | -19.5191
11.4943 | -34.5759 | -19.5283
11.5004 | -34.5764 | -19.5372
11.5065 | -34.5768 | -19.5463
11.5126 | -34.5772 | -19.5553
11.5187 | -34.5777 | -19.5644
11.5248 | -34.5781 | -19.5737
11.5309 | -34.5786 | -19.5827
11.5371 | -34.579 | -19.5917
11.5432 | -34.5794 | -19.6009
11.5493 | -34.5799 | -19.6102
11.5554 | -34.5803 | -19.6193
11.5615 | -34.5807 | -19.6286
11.5676 | -34.5812 | -19.6377
11.5737 | -34.5816 | -19.6466
11.5798 | -34.582 | -19.6555
11.5859 | -34.5825 | -19.6646
11.592 | -34.5829 | -19.6736
11.5981 | -34.5833 | -19.6827
11.6042 | -34.5838 | -19.6917
11.6103 | -34.5842 | -19.7007
11.6164 | -34.5846 | -19.7097
11.6225 | -34.5851 | -19.7187
11.6286 | -34.5855 | -19.7278
11.6347 | -34.5859 | -19.7369
11.6408 | -34.5864 | -19.7459
11.6469 | -34.5868 | -19.755
11.653 | -34.5873 | -19.7639
11.6591 | -34.5877 | -19.7729
11.6652 | -34.5881 | -19.7819
11.6713 | -34.5886 | -19.7909
11.6775 | -34.5891 | -19.7998
11.6836 | -34.5895 | -19.8089
11.6897 | -34.59 | -19.8179
11.6958 | -34.5904 | -19.8271
11.7019 | -34.5908 | -19.8363
11.708 | -34.5913 | -19.8454
11.7141 | -34.5917 | -19.8542
11.7202 | -34.5922 | -19.8632
11.7263 | -34.5926 | -19.8723
11.7324 | -34.593 | -19.8815
11.7385 | -34.5935 | -19.8905
11.7446 | -34.5939 | -19.8995
11.7507 | -34.5944 | -19.9083
11.7568 | -34.5948 | -19.9173
11.7629 | -34.5952 | -19.9263
11.769 | -34.5957 | -19.9353
11.7751 | -34.5961 | -19.9444
11.7812 | -34.5966 | -19.9532
11.7873 | -34.597 | -19.9623
11.7934 | -34.5975 | -19.9714
11.7995 | -34.5979 | -19.9803
11.8056 | -34.5983 | -19.9892
11.8117 | -34.5988 | -19.9982
11.8178 | -34.5992 | -20.0071
11.824 | -34.5997 | -20.0161
11.8301 | -34.6001 | -20.025
11.8362 | -34.6006 | -20.0339
11.8423 | -34.601 | -20.0429
11.8484 | -34.6014 | -20.0519
11.8545 | -34.6019 | -20.0607
11.8606 | -34.6023 | -20.0696
11.8667 | -34.6028 | -20.0786
11.8728 | -34.6032 | -20.0875
11.8789 | -34.6037 | -20.0965
11.885 | -34.6041 | -20.1054
11.8911 | -34.6046 | -20.1143
11.8972 | -34.605 | -20.1233
11.9033 | -34.6055 | -20.1322
11.9094 | -34.6059 | -20.1412
11.9155 | -34.6063 | -20.1501
11.9216 | -34.6068 | -20.1589
11.9277 | -34.6072 | -20.1678
11.9338 | -34.6077 | -20.1768
11.9399 | -34.6081 | -20.1859
11.946 | -34.6086 | -20.1949
11.9521 | -34.609 | -20.2037
11.9582 | -34.6095 | -20.2125
11.9644 | -34.6099 | -20.2214
11.9705 | -34.6104 | -20.2304
11.9766 | -34.6108 | -20.2393
11.9827 | -34.6113 | -20.2483
11.9888 | -34.6118 | -20.2572
11.9949 | -34.6122 | -20.2658
12.001 | -34.6127 | -20.2748
12.0071 | -34.6131 | -20.2836
12.0132 | -34.6136 | -20.2925
12.0193 | -34.614 | -20.3015
12.0254 | -34.6144 | -20.3104
12.0315 | -34.6149 | -20.3191
12.0376 | -34.6153 | -20.3279
12.0437 | -34.6158 | -20.3368
12.0498 | -34.6162 | -20.3458
12.0559 | -34.6167 | -20.3547
12.062 | -34.6172 | -20.3633
12.0681 | -34.6176 | -20.3722
12.0742 | -34.6181 | -20.381
12.0803 | -34.6185 | -20.39
12.0864 | -34.619 | -20.399
12.0925 | -34.6194 | -20.4079
12.0986 | -34.6199 | -20.4167
12.1047 | -34.6203 | -20.4255
12.1109 | -34.6208 | -20.4343
12.117 | -34.6212 | -20.443
12.1231 | -34.6217 | -20.4518
12.1292 | -34.6222 | -20.4605
12.1353 | -34.6226 | -20.4692
12.1414 | -34.6231 | -20.4781
12.1475 | -34.6235 | -20.4871
12.1536 | -34.624 | -20.496
12.1597 | -34.6244 | -20.5049
12.1658 | -34.6249 | -20.5136
12.1719 | -34.6253 | -20.5222
12.178 | -34.6258 | -20.5311
12.1841 | -34.6263 | -20.5401
12.1902 | -34.6267 | -20.5489
12.1963 | -34.6272 | -20.5578
12.2024 | -34.6277 | -20.5664
12.2085 | -34.6281 | -20.5753
12.2146 | -34.6286 | -20.5842
12.2207 | -34.629 | -20.5929
12.2268 | -34.6295 | -20.6017
12.2329 | -34.6299 | -20.6105
12.239 | -34.6304 | -20.619
12.2451 | -34.6309 | -20.6278
12.2513 | -34.6313 | -20.6368
12.2574 | -34.6318 | -20.6456
12.2635 | -34.6322 | -20.6544
12.2696 | -34.6327 | -20.6631
12.2757 | -34.6332 | -20.6716
12.2818 | -34.6337 | -20.6804
12.2879 | -34.6341 | -20.6893
12.294 | -34.6346 | -20.6979
12.3001 | -34.635 | -20.7067
12.3062 | -34.6355 | -20.7155
12.3123 | -34.6359 | -20.724
12.3184 | -34.6364 | -20.7326
12.3245 | -34.6369 | -20.7413
12.3306 | -34.6373 | -20.7501
12.3367 | -34.6378 | -20.7589
12.3428 | -34.6382 | -20.7677
12.3489 | -34.6387 | -20.7761
12.355 | -34.6391 | -20.785
12.3611 | -34.6396 | -20.7939
12.3672 | -34.64 | -20.8026
12.3733 | -34.6405 | -20.8114
12.3794 | -34.641 | -20.8198
12.3855 | -34.6414 | -20.8285
12.3916 | -34.6419 | -20.8372
12.3978 | -34.6424 | -20.846
12.4039 | -34.6428 | -20.8548
12.41 | -34.6433 | -20.8636
12.4161 | -34.6438 | -20.872
12.4222 | -34.6442 | -20.8808
12.4283 | -34.6447 | -20.8894
12.4344 | -34.6452 | -20.8982
12.4405 | -34.6456 | -20.907
12.4466 | -34.6461 | -20.9156
12.4527 | -34.6465 | -20.9243
12.4588 | -34.647 | -20.933
12.4649 | -34.6475 | -20.9417
12.471 | -34.6479 | -20.9505
12.4771 | -34.6484 | -20.9591
12.4832 | -34.6488 | -20.9678
12.4893 | -34.6493 | -20.9763
12.4954 | -34.6498 | -20.9848
12.5015 | -34.6502 | -20.9935
12.5076 | -34.6506 | -21.0024
12.5137 | -34.6511 | -21.011
12.5198 | -34.6516 | -21.0194
12.5259 | -34.6521 | -21.0282
12.532 | -34.6525 | -21.0369
12.5382 | -34.653 | -21.0456
12.5443 | -34.6534 | -21.0543
12.5504 | -34.6539 | -21.063
12.5565 | -34.6544 | -21.0715
12.5626 | -34.6548 | -21.0803
12.5687 | -34.6553 | -21.089
12.5748 | -34.6557 | -21.0977
12.5809 | -34.6562 | -21.1064
12.587 | -34.6567 | -21.1152
12.5931 | -34.6572 | -21.1237
12.5992 | -34.6576 | -21.1324
12.6053 | -34.6581 | -21.141
12.6114 | -34.6585 | -21.1495
12.6175 | -34.659 | -21.1583
12.6236 | -34.6595 | -21.167
12.6297 | -34.6599 | -21.1756
12.6358 | -34.6604 | -21.1841
12.6419 | -34.6609 | -21.1928
12.648 | -34.6613 | -21.2015
12.6541 | -34.6618 | -21.2101
12.6602 | -34.6623 | -21.2185
12.6663 | -34.6627 | -21.2272
12.6724 | -34.6632 | -21.2357
12.6785 | -34.6637 | -21.2443
12.6847 | -34.6641 | -21.2527
12.6908 | -34.6646 | -21.2615
12.6969 | -34.6651 | -21.2699
12.703 | -34.6655 | -21.2786
12.7091 | -34.666 | -21.2873
12.7152 | -34.6664 | -21.2958
12.7213 | -34.6669 | -21.3045
12.7274 | -34.6674 | -21.3131
12.7335 | -34.6678 | -21.3214
12.7396 | -34.6683 | -21.33
12.7457 | -34.6687 | -21.3387
12.7518 | -34.6692 | -21.3473
12.7579 | -34.6697 | -21.356
12.764 | -34.6701 | -21.3647
12.7701 | -34.6706 | -21.3733
12.7762 | -34.671 | -21.3818
12.7823 | -34.6715 | -21.3903
12.7884 | -34.672 | -21.3991
12.7945 | -34.6725 | -21.4078
12.8006 | -34.6729 | -21.4163
12.8067 | -34.6734 | -21.4248
12.8128 | -34.6738 | -21.4334
12.8189 | -34.6743 | -21.4419
12.8251 | -34.6748 | -21.4506
12.8312 | -34.6752 | -21.4593
12.8373 | -34.6757 | -21.4677
12.8434 | -34.6762 | -21.4763
12.8495 | -34.6766 | -21.4848
12.8556 | -34.6771 | -21.4933
12.8617 | -34.6775 | -21.5019
12.8678 | -34.678 | -21.5106
12.8739 | -34.6785 | -21.5189
12.88 | -34.6789 | -21.5275
12.8861 | -34.6794 | -21.536
12.8922 | -34.6798 | -21.5446
12.8983 | -34.6803 | -21.5533
12.9044 | -34.6808 | -21.5618
12.9105 | -34.6812 | -21.5702
12.9166 | -34.6817 | -21.5787
12.9227 | -34.6822 | -21.5873
12.9288 | -34.6826 | -21.5959
12.9349 | -34.6831 | -21.6044
12.941 | -34.6836 | -21.6129
12.9471 | -34.6841 | -21.6215
12.9532 | -34.6845 | -21.63
12.9593 | -34.685 | -21.6387
12.9654 | -34.6855 | -21.6474
12.9716 | -34.6859 | -21.6561
12.9777 | -34.6864 | -21.6647
12.9838 | -34.6869 | -21.6732
12.9899 | -34.6873 | -21.6818
12.996 | -34.6878 | -21.6903
13.0021 | -34.6883 | -21.6989
13.0082 | -34.6888 | -21.7074
13.0143 | -34.6892 | -21.7156
13.0204 | -34.6897 | -21.7241
13.0265 | -34.6902 | -21.7326
13.0326 | -34.6906 | -21.7412
13.0387 | -34.6911 | -21.7499
13.0448 | -34.6915 | -21.7584
13.0509 | -34.692 | -21.7669
13.057 | -34.6925 | -21.7753
13.0631 | -34.693 | -21.7837
13.0692 | -34.6935 | -21.7922
13.0753 | -34.694 | -21.8009
13.0814 | -34.6944 | -21.8095
13.0875 | -34.6949 | -21.8181
13.0936 | -34.6954 | -21.8267
13.0997 | -34.6958 | -21.8351
13.1058 | -34.6963 | -21.8436
13.112 | -34.6968 | -21.8522
13.1181 | -34.6973 | -21.8605
13.1242 | -34.6977 | -21.869
13.1303 | -34.6982 | -21.8774
13.1364 | -34.6986 | -21.8858
13.1425 | -34.6991 | -21.8944
13.1486 | -34.6995 | -21.9028
13.1547 | -34.7 | -21.9113
13.1608 | -34.7004 | -21.9198
13.1669 | -34.7009 | -21.9283
13.173 | -34.7014 | -21.9367
13.1791 | -34.7019 | -21.9452
13.1852 | -34.7024 | -21.9536
13.1913 | -34.7028 | -21.962
13.1974 | -34.7033 | -21.9706
13.2035 | -34.7037 | -21.979
13.2096 | -34.7042 | -21.9873
13.2157 | -34.7047 | -21.9957
13.2218 | -34.7051 | -22.0042
13.2279 | -34.7056 | -22.0127
13.234 | -34.7061 | -22.0211
13.2401 | -34.7066 | -22.0296
13.2462 | -34.707 | -22.0381
13.2524 | -34.7075 | -22.0467
13.2585 | -34.708 | -22.0552
13.2646 | -34.7084 | -22.0637
13.2707 | -34.7089 | -22.0722
13.2768 | -34.7094 | -22.0809
13.2829 | -34.7099 | -22.0894
13.289 | -34.7103 | -22.0977
13.2951 | -34.7108 | -22.1062
13.3012 | -34.7113 | -22.1147
13.3073 | -34.7117 | -22.1234
13.3134 | -34.7122 | -22.1319
13.3195 | -34.7127 | -22.1402
13.3256 | -34.7131 | -22.1486
13.3317 | -34.7136 | -22.1568
13.3378 | -34.7141 | -22.1654
13.3439 | -34.7146 | -22.174
13.35 | -34.7151 | -22.1825
13.3561 | -34.7155 | -22.1908
13.3622 | -34.716 | -22.1992
13.3683 | -34.7164 | -22.2076
13.3744 | -34.7169 | -22.2159
13.3805 | -34.7174 | -22.2246
13.3866 | -34.7178 | -22.2329
13.3927 | -34.7183 | -22.2414
13.3989 | -34.7187 | -22.2498
13.405 | -34.7192 | -22.2583
13.4111 | -34.7196 | -22.2666
13.4172 | -34.7201 | -22.2752
13.4233 | -34.7206 | -22.2837
13.4294 | -34.721 | -22.2924
13.4355 | -34.7215 | -22.3007
13.4416 | -34.7219 | -22.309
13.4477 | -34.7224 | -22.3174
13.4538 | -34.7229 | -22.3258
13.4599 | -34.7234 | -22.3343
13.466 | -34.7238 | -22.3429
13.4721 | -34.7243 | -22.3513
13.4782 | -34.7248 | -22.3598
13.4843 | -34.7252 | -22.3684
13.4904 | -34.7257 | -22.3768
13.4965 | -34.7262 | -22.3852
13.5026 | -34.7266 | -22.3935
13.5087 | -34.7271 | -22.4019
13.5148 | -34.7275 | -22.4104
13.5209 | -34.728 | -22.4189
13.527 | -34.7284 | -22.4274
13.5331 | -34.7289 | -22.4358
13.5393 | -34.7293 | -22.4441
13.5454 | -34.7298 | -22.4524
13.5515 | -34.7302 | -22.4611
13.5576 | -34.7307 | -22.4696
13.5637 | -34.7312 | -22.4782
13.5698 | -34.7316 | -22.4868
13.5759 | -34.7321 | -22.495
13.582 | -34.7326 | -22.5032
13.5881 | -34.7331 | -22.5115
13.5942 | -34.7336 | -22.52
13.6003 | -34.7341 | -22.5284
13.6064 | -34.7346 | -22.5368
13.6125 | -34.735 | -22.5451
13.6186 | -34.7355 | -22.5535
13.6247 | -34.736 | -22.562
13.6308 | -34.7365 | -22.5703
13.6369 | -34.7369 | -22.5789
13.643 | -34.7374 | -22.5874
13.6491 | -34.7378 | -22.5957
13.6552 | -34.7383 | -22.6041
13.6613 | -34.7388 | -22.6125
13.6674 | -34.7393 | -22.621
13.6735 | -34.7397 | -22.6294
13.6796 | -34.7402 | -22.6377
13.6858 | -34.7407 | -22.6459
13.6919 | -34.7411 | -22.6545
13.698 | -34.7416 | -22.6629
13.7041 | -34.7421 | -22.6715
13.7102 | -34.7426 | -22.6799
13.7163 | -34.7431 | -22.6883
13.7224 | -34.7435 | -22.6969
13.7285 | -34.744 | -22.7052
13.7346 | -34.7444 | -22.7137
13.7407 | -34.7449 | -22.722
13.7468 | -34.7454 | -22.7303
13.7529 | -34.7458 | -22.7388
13.759 | -34.7463 | -22.7475
13.7651 | -34.7467 | -22.7558
13.7712 | -34.7472 | -22.7643
13.7773 | -34.7477 | -22.773
13.7834 | -34.7481 | -22.7814
13.7895 | -34.7486 | -22.7898
13.7956 | -34.7491 | -22.7984
13.8017 | -34.7495 | -22.8067
13.8078 | -34.75 | -22.8151
13.8139 | -34.7505 | -22.8235
13.82 | -34.7509 | -22.8319
13.8262 | -34.7514 | -22.8404
13.8323 | -34.7519 | -22.8489
13.8384 | -34.7524 | -22.8575
13.8445 | -34.7529 | -22.8659
13.8506 | -34.7533 | -22.8743
13.8567 | -34.7538 | -22.8829
13.8628 | -34.7543 | -22.8912
13.8689 | -34.7547 | -22.8995
13.875 | -34.7552 | -22.9077
13.8811 | -34.7557 | -22.9163
13.8872 | -34.7561 | -22.9247
13.8933 | -34.7565 | -22.9331
13.8994 | -34.757 | -22.9416
13.9055 | -34.7575 | -22.95
13.9116 | -34.7579 | -22.9583
13.9177 | -34.7584 | -22.9667
13.9238 | -34.7588 | -22.9752
13.9299 | -34.7593 | -22.9833
13.936 | -34.7598 | -22.9919
13.9421 | -34.7602 | -23.0003
13.9482 | -34.7607 | -23.0086
13.9543 | -34.7612 | -23.0173
13.9604 | -34.7616 | -23.0256
13.9665 | -34.7621 | -23.0341
13.9727 | -34.7626 | -23.0428
13.9788 | -34.7631 | -23.0512
13.9849 | -34.7635 | -23.0596
13.991 | -34.764 | -23.0681
13.9971 | -34.7645 | -23.0762
14.0032 | -34.7649 | -23.0845
14.0093 | -34.7654 | -23.0931
14.0154 | -34.7658 | -23.1014
14.0215 | -34.7663 | -23.1098
14.0276 | -34.7668 | -23.1181
14.0337 | -34.7672 | -23.1265
14.0398 | -34.7677 | -23.1347
14.0459 | -34.7681 | -23.1431
14.052 | -34.7686 | -23.1515
14.0581 | -34.769 | -23.1601
14.0642 | -34.7695 | -23.1685
14.0703 | -34.77 | -23.1771
14.0764 | -34.7704 | -23.1854
14.0825 | -34.7709 | -23.1938
14.0886 | -34.7713 | -23.2022
14.0947 | -34.7718 | -23.2105
14.1008 | -34.7723 | -23.219
14.1069 | -34.7727 | -23.2272
14.1131 | -34.7732 | -23.2356
14.1192 | -34.7736 | -23.2439
14.1253 | -34.7741 | -23.2522
14.1314 | -34.7746 | -23.2605
14.1375 | -34.7751 | -23.2691
14.1436 | -34.7755 | -23.2774
14.1497 | -34.776 | -23.2857
14.1558 | -34.7765 | -23.2941
14.1619 | -34.777 | -23.3024
14.168 | -34.7774 | -23.3108
14.1741 | -34.7779 | -23.319
14.1802 | -34.7784 | -23.3274
14.1863 | -34.7789 | -23.3359
14.1924 | -34.7793 | -23.3443
14.1985 | -34.7798 | -23.3527
14.2046 | -34.7803 | -23.3613
14.2107 | -34.7808 | -23.3694
14.2168 | -34.7812 | -23.3778
14.2229 | -34.7816 | -23.3863
14.229 | -34.7821 | -23.3948
14.2351 | -34.7826 | -23.4031
14.2412 | -34.783 | -23.4115
14.2473 | -34.7835 | -23.4197
14.2534 | -34.784 | -23.428
14.2596 | -34.7844 | -23.4362
14.2657 | -34.7849 | -23.4446
14.2718 | -34.7854 | -23.4531
14.2779 | -34.7858 | -23.4614
14.284 | -34.7863 | -23.4698
14.2901 | -34.7868 | -23.4781
14.2962 | -34.7873 | -23.4865
14.3023 | -34.7877 | -23.4948
14.3084 | -34.7882 | -23.5033
14.3145 | -34.7887 | -23.5116
14.3206 | -34.7891 | -23.5202
14.3267 | -34.7896 | -23.5286
14.3328 | -34.7901 | -23.537
14.3389 | -34.7905 | -23.5453
14.345 | -34.791 | -23.5537
14.3511 | -34.7914 | -23.5622
14.3572 | -34.7919 | -23.5707
14.3633 | -34.7924 | -23.579
14.3694 | -34.7928 | -23.5873
14.3755 | -34.7933 | -23.5956
14.3816 | -34.7938 | -23.6041
14.3877 | -34.7943 | -23.6122
14.3938 | -34.7947 | -23.6207
14.4 | -34.7952 | -23.629
14.4061 | -34.7957 | -23.6374
14.4122 | -34.7961 | -23.6458
14.4183 | -34.7966 | -23.6541
14.4244 | -34.797 | -23.6627
14.4305 | -34.7975 | -23.6709
14.4366 | -34.798 | -23.6793
14.4427 | -34.7985 | -23.6879
14.4488 | -34.7989 | -23.6963
14.4549 | -34.7994 | -23.7047
14.461 | -34.7999 | -23.7128
14.4671 | -34.8003 | -23.7213
14.4732 | -34.8008 | -23.7296
14.4793 | -34.8013 | -23.738
14.4854 | -34.8018 | -23.7462
14.4915 | -34.8022 | -23.7544
14.4976 | -34.8027 | -23.7626
14.5037 | -34.8032 | -23.771
14.5098 | -34.8037 | -23.7793
14.5159 | -34.8041 | -23.7876
14.522 | -34.8046 | -23.7958
14.5281 | -34.805 | -23.8039
14.5342 | -34.8055 | -23.8123
14.5403 | -34.806 | -23.8205
14.5465 | -34.8065 | -23.829
14.5526 | -34.8069 | -23.8373
14.5587 | -34.8074 | -23.8457
14.5648 | -34.8079 | -23.8541
14.5709 | -34.8084 | -23.8625
14.577 | -34.8088 | -23.8707
14.5831 | -34.8092 | -23.879
14.5892 | -34.8097 | -23.8874
14.5953 | -34.8102 | -23.8957
14.6014 | -34.8107 | -23.904
14.6075 | -34.8111 | -23.9123
14.6136 | -34.8116 | -23.9206
14.6197 | -34.812 | -23.929
14.6258 | -34.8125 | -23.9373
14.6319 | -34.8129 | -23.9456
14.638 | -34.8134 | -23.9538
14.6441 | -34.8138 | -23.9622
14.6502 | -34.8143 | -23.9706
14.6563 | -34.8148 | -23.9788
14.6624 | -34.8153 | -23.9872
14.6685 | -34.8158 | -23.9955
14.6746 | -34.8163 | -24.0039
14.6807 | -34.8167 | -24.0121
14.6869 | -34.8172 | -24.0205
14.693 | -34.8177 | -24.0287
14.6991 | -34.8181 | -24.037
14.7052 | -34.8186 | -24.0453
14.7113 | -34.8191 | -24.0536
14.7174 | -34.8196 | -24.0619
14.7235 | -34.82 | -24.0701
14.7296 | -34.8205 | -24.0784
14.7357 | -34.821 | -24.0866
14.7418 | -34.8214 | -24.0949
14.7479 | -34.8219 | -24.1032
14.754 | -34.8224 | -24.1116
14.7601 | -34.8228 | -24.1198
14.7662 | -34.8233 | -24.128
14.7723 | -34.8238 | -24.136
14.7784 | -34.8242 | -24.1443
14.7845 | -34.8247 | -24.1526
14.7906 | -34.8252 | -24.1609
14.7967 | -34.8257 | -24.1691
14.8028 | -34.8261 | -24.1773
14.8089 | -34.8266 | -24.1856
14.815 | -34.8271 | -24.1938
14.8211 | -34.8275 | -24.202
14.8272 | -34.828 | -24.2104
14.8334 | -34.8284 | -24.2186
14.8395 | -34.8289 | -24.2269
14.8456 | -34.8294 | -24.2352
14.8517 | -34.8299 | -24.2436
14.8578 | -34.8303 | -24.252
14.8639 | -34.8308 | -24.2604
14.87 | -34.8313 | -24.2687
14.8761 | -34.8318 | -24.2767
14.8822 | -34.8322 | -24.285
14.8883 | -34.8327 | -24.2934
14.8944 | -34.8332 | -24.3015
14.9005 | -34.8336 | -24.3098
14.9066 | -34.8341 | -24.3181
14.9127 | -34.8346 | -24.3262
14.9188 | -34.835 | -24.3344
14.9249 | -34.8355 | -24.3427
14.931 | -34.836 | -24.351
14.9371 | -34.8365 | -24.3593
14.9432 | -34.8369 | -24.3676
14.9493 | -34.8374 | -24.3758
14.9554 | -34.8379 | -24.384
14.9615 | -34.8383 | -24.3923
14.9676 | -34.8388 | -24.4007
14.9738 | -34.8392 | -24.4088
14.9799 | -34.8397 | -24.4173
14.986 | -34.8402 | -24.4254
14.9921 | -34.8406 | -24.4337
14.9982 | -34.8411 | -24.4419
15.0043 | -34.8416 | -24.4501
15.0104 | -34.842 | -24.4584
15.0165 | -34.8425 | -24.4668
15.0226 | -34.843 | -24.4749
15.0287 | -34.8435 | -24.4833
15.0348 | -34.8439 | -24.4916
15.0409 | -34.8444 | -24.4999
15.047 | -34.8449 | -24.5083
15.0531 | -34.8454 | -24.5164
15.0592 | -34.8458 | -24.5246
15.0653 | -34.8463 | -24.533
15.0714 | -34.8468 | -24.5411
15.0775 | -34.8472 | -24.5492
15.0836 | -34.8477 | -24.5575
15.0897 | -34.8482 | -24.5654
15.0958 | -34.8487 | -24.5736
15.1019 | -34.8491 | -24.5819
15.108 | -34.8496 | -24.5901
15.1141 | -34.8501 | -24.5983
15.1203 | -34.8506 | -24.6067
15.1264 | -34.8511 | -24.615
15.1325 | -34.8515 | -24.6232
15.1386 | -34.852 | -24.6314
15.1447 | -34.8525 | -24.6396
15.1508 | -34.853 | -24.6478
15.1569 | -34.8535 | -24.656
15.163 | -34.8539 | -24.664
15.1691 | -34.8544 | -24.672
15.1752 | -34.8549 | -24.6803
15.1813 | -34.8553 | -24.6884
15.1874 | -34.8558 | -24.6967
15.1935 | -34.8563 | -24.7046
15.1996 | -34.8567 | -24.7127
15.2057 | -34.8572 | -24.7208
15.2118 | -34.8577 | -24.729
15.2179 | -34.8582 | -24.7372
15.224 | -34.8586 | -24.7454
15.2301 | -34.8591 | -24.7533
15.2362 | -34.8596 | -24.7615
15.2423 | -34.8601 | -24.7698
15.2484 | -34.8605 | -24.778
15.2545 | -34.861 | -24.7862
15.2607 | -34.8615 | -24.7943
15.2668 | -34.862 | -24.8022
15.2729 | -34.8625 | -24.8103
15.279 | -34.8629 | -24.8186
15.2851 | -34.8634 | -24.8266
15.2912 | -34.8639 | -24.8347
15.2973 | -34.8644 | -24.843
15.3034 | -34.8649 | -24.851
15.3095 | -34.8653 | -24.8593
15.3156 | -34.8658 | -24.8674
15.3217 | -34.8663 | -24.8754
15.3278 | -34.8667 | -24.8837
15.3339 | -34.8672 | -24.8918
15.34 | -34.8677 | -24.8999
15.3461 | -34.8682 | -24.908
15.3522 | -34.8686 | -24.9162
15.3583 | -34.8691 | -24.9245
15.3644 | -34.8695 | -24.9326
15.3705 | -34.87 | -24.9405
15.3766 | -34.8705 | -24.9487
15.3827 | -34.871 | -24.9567
15.3888 | -34.8715 | -24.965
15.3949 | -34.8719 | -24.9732
15.401 | -34.8724 | -24.9815
15.4072 | -34.8728 | -24.9893
15.4133 | -34.8733 | -24.9975
15.4194 | -34.8738 | -25.0058
15.4255 | -34.8742 | -25.0139
15.4316 | -34.8747 | -25.0221
15.4377 | -34.8752 | -25.0302
15.4438 | -34.8757 | -25.0382
15.4499 | -34.8762 | -25.0464
15.456 | -34.8767 | -25.0546
15.4621 | -34.8771 | -25.0626
15.4682 | -34.8776 | -25.0709
15.4743 | -34.8781 | -25.0787
15.4804 | -34.8786 | -25.0868
15.4865 | -34.8791 | -25.095
15.4926 | -34.8795 | -25.1032
15.4987 | -34.88 | -25.1113
15.5048 | -34.8805 | -25.1193
15.5109 | -34.881 | -25.1272
15.517 | -34.8815 | -25.1352
15.5231 | -34.8819 | -25.1434
15.5292 | -34.8824 | -25.1514
15.5353 | -34.8828 | -25.1596
15.5414 | -34.8833 | -25.1676
15.5476 | -34.8838 | -25.1754
15.5537 | -34.8843 | -25.1835
15.5598 | -34.8848 | -25.1917
15.5659 | -34.8852 | -25.1998
15.572 | -34.8857 | -25.2078
15.5781 | -34.8862 | -25.2158
15.5842 | -34.8867 | -25.2236
15.5903 | -34.8872 | -25.2316
15.5964 | -34.8876 | -25.2397
15.6025 | -34.8881 | -25.2478
15.6086 | -34.8886 | -25.2559
15.6147 | -34.889 | -25.264
15.6208 | -34.8895 | -25.2719
15.6269 | -34.89 | -25.2799
15.633 | -34.8905 | -25.288
15.6391 | -34.891 | -25.2962
15.6452 | -34.8914 | -25.3043
15.6513 | -34.892 | -25.3121
15.6574 | -34.8924 | -25.3203
15.6635 | -34.8929 | -25.3283
15.6696 | -34.8934 | -25.3364
15.6757 | -34.8938 | -25.3445
15.6818 | -34.8943 | -25.3526
15.688 | -34.8948 | -25.3605
15.6941 | -34.8953 | -25.3685
15.7002 | -34.8958 | -25.3765
15.7063 | -34.8962 | -25.3845
15.7124 | -34.8967 | -25.3927
15.7185 | -34.8972 | -25.4007
15.7246 | -34.8977 | -25.4085
15.7307 | -34.8981 | -25.4165
15.7368 | -34.8986 | -25.4245
15.7429 | -34.8991 | -25.4326
15.749 | -34.8996 | -25.4407
15.7551 | -34.9 | -25.4489
15.7612 | -34.9005 | -25.4569
15.7673 | -34.901 | -25.4649
15.7734 | -34.9015 | -25.473
15.7795 | -34.9019 | -25.4809
15.7856 | -34.9024 | -25.4889
15.7917 | -34.9029 | -25.4966
15.7978 | -34.9034 | -25.5048
15.8039 | -34.9039 | -25.5127
15.81 | -34.9044 | -25.5207
15.8161 | -34.9048 | -25.5286
15.8222 | -34.9053 | -25.5365
15.8283 | -34.9058 | -25.5443
15.8345 | -34.9063 | -25.5523
15.8406 | -34.9067 | -25.5603
15.8467 | -34.9072 | -25.5685
15.8528 | -34.9077 | -25.5764
15.8589 | -34.9082 | -25.5844
15.865 | -34.9087 | -25.592
15.8711 | -34.9092 | -25.6002
15.8772 | -34.9097 | -25.608
15.8833 | -34.9101 | -25.6161
15.8894 | -34.9106 | -25.6242
15.8955 | -34.9111 | -25.6322
15.9016 | -34.9116 | -25.6398
15.9077 | -34.912 | -25.6477
15.9138 | -34.9125 | -25.6556
15.9199 | -34.913 | -25.6635
15.926 | -34.9134 | -25.6715
15.9321 | -34.9139 | -25.6791
15.9382 | -34.9144 | -25.687
15.9443 | -34.9149 | -25.695
15.9504 | -34.9153 | -25.7031
15.9565 | -34.9158 | -25.711
15.9626 | -34.9163 | -25.719
15.9687 | -34.9168 | -25.7267
15.9749 | -34.9172 | -25.7345
15.981 | -34.9177 | -25.7425
15.9871 | -34.9182 | -25.7506
15.9932 | -34.9187 | -25.7584
15.9993 | -34.9191 | -25.7664
16.0054 | -34.9196 | -25.774
16.0115 | -34.9201 | -25.782
16.0176 | -34.9206 | -25.7901
16.0237 | -34.921 | -25.7981
16.0298 | -34.9215 | -25.8059
16.0359 | -34.922 | -25.8138
16.042 | -34.9225 | -25.8214
16.0481 | -34.9229 | -25.8293
16.0542 | -34.9234 | -25.8373
16.0603 | -34.9238 | -25.8451
16.0664 | -34.9243 | -25.8529
16.0725 | -34.9248 | -25.8607
16.0786 | -34.9253 | -25.8686
16.0847 | -34.9258 | -25.8765
16.0908 | -34.9262 | -25.8846
16.0969 | -34.9267 | -25.8925
16.103 | -34.9272 | -25.9004
16.1091 | -34.9277 | -25.9082
16.1152 | -34.9282 | -25.916
16.1214 | -34.9286 | -25.9239
16.1275 | -34.9291 | -25.9317
16.1336 | -34.9296 | -25.9394
16.1397 | -34.9301 | -25.9472
16.1458 | -34.9306 | -25.9551
16.1519 | -34.9311 | -25.9631
16.158 | -34.9315 | -25.9711
16.1641 | -34.932 | -25.9792
16.1702 | -34.9325 | -25.9871
16.1763 | -34.9329 | -25.995
16.1824 | -34.9334 | -26.0027
16.1885 | -34.9339 | -26.0106
16.1946 | -34.9344 | -26.0185
16.2007 | -34.9349 | -26.0264
16.2068 | -34.9353 | -26.0343
16.2129 | -34.9358 | -26.0422
16.219 | -34.9363 | -26.0498
16.2251 | -34.9368 | -26.0575
16.2312 | -34.9373 | -26.0653
16.2373 | -34.9377 | -26.073
16.2434 | -34.9382 | -26.0809
16.2495 | -34.9387 | -26.0886
16.2556 | -34.9392 | -26.0965
16.2618 | -34.9397 | -26.1043
16.2679 | -34.9401 | -26.1122
16.274 | -34.9406 | -26.1201
16.2801 | -34.9411 | -26.1278
16.2862 | -34.9416 | -26.1353
16.2923 | -34.9421 | -26.1432
16.2984 | -34.9425 | -26.1511
16.3045 | -34.943 | -26.159
16.3106 | -34.9434 | -26.1669
16.3167 | -34.9439 | -26.1749
16.3228 | -34.9444 | -26.1824
16.3289 | -34.9448 | -26.1902
16.335 | -34.9453 | -26.198
16.3411 | -34.9458 | -26.2058
16.3472 | -34.9463 | -26.2137
16.3533 | -34.9468 | -26.2214
16.3594 | -34.9472 | -26.2289
16.3655 | -34.9477 | -26.2367
16.3716 | -34.9481 | -26.2445
16.3777 | -34.9486 | -26.2522
16.3838 | -34.9491 | -26.2601
16.3899 | -34.9496 | -26.2678
16.396 | -34.9501 | -26.2757
16.4021 | -34.9506 | -26.2834
16.4083 | -34.951 | -26.2913
16.4144 | -34.9515 | -26.2991
16.4205 | -34.952 | -26.3069
16.4266 | -34.9525 | -26.3144
16.4327 | -34.9529 | -26.3221
16.4388 | -34.9534 | -26.33
16.4449 | -34.9539 | -26.3379
16.451 | -34.9544 | -26.3457
16.4571 | -34.9548 | -26.3535
16.4632 | -34.9553 | -26.3611
16.4693 | -34.9557 | -26.3689
16.4754 | -34.9562 | -26.3769
16.4815 | -34.9567 | -26.3846
16.4876 | -34.9572 | -26.3923
16.4937 | -34.9576 | -26.4001
16.4998 | -34.9581 | -26.4077
16.5059 | -34.9586 | -26.4156
16.512 | -34.9591 | -26.4234
16.5181 | -34.9596 | -26.4312
16.5242 | -34.9601 | -26.439
16.5303 | -34.9605 | -26.4467
16.5364 | -34.961 | -26.4542
16.5425 | -34.9615 | -26.462
16.5487 | -34.9619 | -26.4697
16.5548 | -34.9624 | -26.4776
16.5609 | -34.9629 | -26.4853
16.567 | -34.9634 | -26.493
16.5731 | -34.9638 | -26.5009
16.5792 | -34.9643 | -26.5088
16.5853 | -34.9648 | -26.5166
16.5914 | -34.9652 | -26.5245
16.5975 | -34.9657 | -26.5324
16.6036 | -34.9662 | -26.5398
16.6097 | -34.9666 | -26.5476
16.6158 | -34.9671 | -26.5553
16.6219 | -34.9676 | -26.5632
16.628 | -34.968 | -26.5709
16.6341 | -34.9685 | -26.5786
16.6402 | -34.969 | -26.5862
16.6463 | -34.9694 | -26.594
16.6524 | -34.9699 | -26.6016
16.6585 | -34.9704 | -26.6095
16.6646 | -34.9709 | -26.6172
16.6707 | -34.9714 | -26.6247
16.6768 | -34.9718 | -26.6325
16.6829 | -34.9723 | -26.6402
16.689 | -34.9728 | -26.648
16.6952 | -34.9732 | -26.6558
16.7013 | -34.9737 | -26.6635
16.7074 | -34.9742 | -26.671
16.7135 | -34.9747 | -26.6788
16.7196 | -34.9751 | -26.6865
16.7257 | -34.9756 | -26.6941
16.7318 | -34.9761 | -26.7018
16.7379 | -34.9766 | -26.7094
16.744 | -34.9771 | -26.717
16.7501 | -34.9775 | -26.7248
16.7562 | -34.978 | -26.7325
16.7623 | -34.9785 | -26.7401
16.7684 | -34.9789 | -26.748
16.7745 | -34.9794 | -26.7557
16.7806 | -34.9799 | -26.7629
16.7867 | -34.9804 | -26.7707
16.7928 | -34.9809 | -26.7783
16.7989 | -34.9813 | -26.7858
16.805 | -34.9818 | -26.7934
16.8111 | -34.9823 | -26.801
16.8172 | -34.9828 | -26.8089
16.8233 | -34.9833 | -26.8166
16.8294 | -34.9837 | -26.8243
16.8356 | -34.9842 | -26.8321
16.8417 | -34.9847 | -26.8397
16.8478 | -34.9852 | -26.8473
16.8539 | -34.9856 | -26.855
16.86 | -34.9861 | -26.8628
16.8661 | -34.9866 | -26.8703
16.8722 | -34.9871 | -26.8781
16.8783 | -34.9876 | -26.8859
16.8844 | -34.9881 | -26.8933
16.8905 | -34.9885 | -26.901
16.8966 | -34.989 | -26.9086
16.9027 | -34.9895 | -26.9163
16.9088 | -34.99 | -26.924
16.9149 | -34.9905 | -26.9317
16.921 | -34.991 | -26.9394
16.9271 | -34.9915 | -26.9471
16.9332 | -34.9919 | -26.9546
16.9393 | -34.9924 | -26.9622
16.9454 | -34.9929 | -26.9697
16.9515 | -34.9934 | -26.9774
16.9576 | -34.9939 | -26.985
16.9637 | -34.9944 | -26.9926
16.9698 | -34.9948 | -27.0001
16.9759 | -34.9953 | -27.0078
16.9821 | -34.9958 | -27.0154
16.9882 | -34.9963 | -27.0229
16.9943 | -34.9968 | -27.0304
17.0004 | -34.9973 | -27.0382
17.0065 | -34.9978 | -27.0459
17.0126 | -34.9982 | -27.0532
17.0187 | -34.9987 | -27.0609
17.0248 | -34.9991 | -27.0685
17.0309 | -34.9996 | -27.0761
17.037 | -35.0001 | -27.0838
17.0431 | -35.0006 | -27.0914
17.0492 | -35.001 | -27.0991
17.0553 | -35.0015 | -27.1068
17.0614 | -35.002 | -27.1143
17.0675 | -35.0024 | -27.1218
17.0736 | -35.0029 | -27.1295
17.0797 | -35.0034 | -27.137
17.0858 | -35.0038 | -27.1448
17.0919 | -35.0043 | -27.1524
17.098 | -35.0048 | -27.16
17.1041 | -35.0053 | -27.1673
17.1102 | -35.0058 | -27.1751
17.1163 | -35.0062 | -27.1824
17.1225 | -35.0067 | -27.1902
17.1286 | -35.0072 | -27.1973
17.1347 | -35.0077 | -27.2049
17.1408 | -35.0082 | -27.2124
17.1469 | -35.0087 | -27.2201
17.153 | -35.0091 | -27.2276
17.1591 | -35.0096 | -27.2352
17.1652 | -35.0101 | -27.2427
17.1713 | -35.0106 | -27.2501
17.1774 | -35.011 | -27.2576
17.1835 | -35.0115 | -27.2652
17.1896 | -35.012 | -27.2728
17.1957 | -35.0125 | -27.2807
17.2018 | -35.013 | -27.288
17.2079 | -35.0135 | -27.2956
17.214 | -35.0139 | -27.3033
17.2201 | -35.0144 | -27.3109
17.2262 | -35.0148 | -27.3185
17.2323 | -35.0153 | -27.3262
17.2384 | -35.0158 | -27.3336
17.2445 | -35.0163 | -27.341
17.2506 | -35.0168 | -27.3488
17.2567 | -35.0172 | -27.3563
17.2628 | -35.0177 | -27.3639
17.269 | -35.0182 | -27.3713
17.2751 | -35.0187 | -27.3788
17.2812 | -35.0191 | -27.3864
17.2873 | -35.0196 | -27.3941
17.2934 | -35.0201 | -27.4018
17.2995 | -35.0206 | -27.4094
17.3056 | -35.0211 | -27.4168
17.3117 | -35.0215 | -27.4244
17.3178 | -35.022 | -27.4319
17.3239 | -35.0225 | -27.4394
17.33 | -35.0229 | -27.447
17.3361 | -35.0234 | -27.4545
17.3422 | -35.0239 | -27.4618
17.3483 | -35.0244 | -27.4693
17.3544 | -35.0249 | -27.4768
17.3605 | -35.0253 | -27.4842
17.3666 | -35.0258 | -27.4917
17.3727 | -35.0263 | -27.4993
17.3788 | -35.0268 | -27.5066
17.3849 | -35.0273 | -27.5141
17.391 | -35.0278 | -27.5217
17.3971 | -35.0282 | -27.5293
17.4032 | -35.0287 | -27.5368
17.4094 | -35.0292 | -27.5443
17.4155 | -35.0297 | -27.5517
17.4216 | -35.0302 | -27.5592
17.4277 | -35.0307 | -27.5668
17.4338 | -35.0311 | -27.5744
17.4399 | -35.0316 | -27.582
17.446 | -35.0321 | -27.5895
17.4521 | -35.0326 | -27.5971
17.4582 | -35.0331 | -27.6044
17.4643 | -35.0336 | -27.6118
17.4704 | -35.034 | -27.6192
17.4765 | -35.0345 | -27.6268
17.4826 | -35.035 | -27.634
17.4887 | -35.0354 | -27.6416
17.4948 | -35.0359 | -27.6493
17.5009 | -35.0364 | -27.6567
17.507 | -35.0369 | -27.6643
17.5131 | -35.0374 | -27.6719
17.5192 | -35.0378 | -27.6791
17.5253 | -35.0383 | -27.6866
17.5314 | -35.0388 | -27.6941
17.5375 | -35.0392 | -27.7016
17.5436 | -35.0397 | -27.7093
17.5497 | -35.0402 | -27.7167
17.5559 | -35.0406 | -27.724
17.562 | -35.0412 | -27.7315
17.5681 | -35.0416 | -27.7389
17.5742 | -35.0421 | -27.7464
17.5803 | -35.0426 | -27.7539
17.5864 | -35.0431 | -27.7612
17.5925 | -35.0436 | -27.7688
17.5986 | -35.0441 | -27.7764
17.6047 | -35.0446 | -27.7839
17.6108 | -35.045 | -27.7913
17.6169 | -35.0455 | -27.7987
17.623 | -35.046 | -27.806
17.6291 | -35.0465 | -27.8134
17.6352 | -35.047 | -27.8207
17.6413 | -35.0474 | -27.8283
17.6474 | -35.0479 | -27.8358
17.6535 | -35.0484 | -27.8432
17.6596 | -35.0489 | -27.8506
17.6657 | -35.0494 | -27.8581
17.6718 | -35.0499 | -27.8656
17.6779 | -35.0504 | -27.8729
17.684 | -35.0509 | -27.8804
17.6901 | -35.0513 | -27.8878
17.6963 | -35.0518 | -27.8952
17.7024 | -35.0523 | -27.9025
17.7085 | -35.0528 | -27.9099
17.7146 | -35.0533 | -27.9172
17.7207 | -35.0537 | -27.9247
17.7268 | -35.0543 | -27.9319
17.7329 | -35.0547 | -27.9394
17.739 | -35.0552 | -27.947
17.7451 | -35.0557 | -27.9544
17.7512 | -35.0562 | -27.9618
17.7573 | -35.0567 | -27.9693
17.7634 | -35.0571 | -27.9767
17.7695 | -35.0576 | -27.984
17.7756 | -35.0581 | -27.9913
17.7817 | -35.0586 | -27.9989
17.7878 | -35.059 | -28.0063
17.7939 | -35.0595 | -28.0139
17.8 | -35.06 | -28.0215
17.8061 | -35.0605 | -28.029
17.8122 | -35.061 | -28.0365
17.8183 | -35.0615 | -28.0438
17.8244 | -35.062 | -28.0512
17.8305 | -35.0624 | -28.0587
17.8366 | -35.0629 | -28.0659
17.8428 | -35.0634 | -28.0734
17.8489 | -35.0638 | -28.0807
17.855 | -35.0643 | -28.0881
17.8611 | -35.0648 | -28.0955
17.8672 | -35.0653 | -28.1027
17.8733 | -35.0658 | -28.1102
17.8794 | -35.0663 | -28.1178
17.8855 | -35.0668 | -28.1253
17.8916 | -35.0673 | -28.1326
17.8977 | -35.0677 | -28.1401
17.9038 | -35.0682 | -28.1474
17.9099 | -35.0687 | -28.155
17.916 | -35.0691 | -28.1625
17.9221 | -35.0696 | -28.1703
17.9282 | -35.0701 | -28.1777
17.9343 | -35.0705 | -28.1853
17.9404 | -35.071 | -28.1924
17.9465 | -35.0715 | -28.1997
17.9526 | -35.072 | -28.2072
17.9587 | -35.0725 | -28.2147
17.9648 | -35.0729 | -28.2219
17.9709 | -35.0734 | -28.2293
17.977 | -35.0739 | -28.2368
17.9832 | -35.0744 | -28.2442
17.9893 | -35.0749 | -28.2515
17.9954 | -35.0754 | -28.259
18.0015 | -35.0759 | -28.2664
18.0076 | -35.0763 | -28.2735
18.0137 | -35.0768 | -28.281
18.0198 | -35.0773 | -28.2882
18.0259 | -35.0778 | -28.2959
18.032 | -35.0783 | -28.3035
18.0381 | -35.0788 | -28.3108
18.0442 | -35.0792 | -28.3181
18.0503 | -35.0797 | -28.3254
18.0564 | -35.0802 | -28.3328
18.0625 | -35.0806 | -28.3401
18.0686 | -35.0811 | -28.3475
18.0747 | -35.0816 | -28.3551
18.0808 | -35.0821 | -28.3624
18.0869 | -35.0826 | -28.3698
18.093 | -35.083 | -28.3771
18.0991 | -35.0835 | -28.3846
18.1052 | -35.084 | -28.3919
18.1113 | -35.0844 | -28.3992
18.1174 | -35.085 | -28.4064
18.1236 | -35.0855 | -28.4137
18.1297 | -35.0859 | -28.4211
18.1358 | -35.0864 | -28.4285
18.1419 | -35.0869 | -28.4359
18.148 | -35.0874 | -28.4433
18.1541 | -35.0879 | -28.4504
18.1602 | -35.0884 | -28.4579
18.1663 | -35.0888 | -28.4652
18.1724 | -35.0893 | -28.4724
18.1785 | -35.0898 | -28.4797
18.1846 | -35.0903 | -28.4869
18.1907 | -35.0908 | -28.4942
18.1968 | -35.0912 | -28.5015
18.2029 | -35.0917 | -28.5088
18.209 | -35.0922 | -28.5162
18.2151 | -35.0927 | -28.5237
18.2212 | -35.0932 | -28.531
18.2273 | -35.0937 | -28.5383
18.2334 | -35.0942 | -28.5455
18.2395 | -35.0946 | -28.5529
18.2456 | -35.0951 | -28.5602
18.2517 | -35.0956 | -28.5678
18.2578 | -35.0961 | -28.5749
18.2639 | -35.0966 | -28.5822
18.2701 | -35.0971 | -28.5895
18.2762 | -35.0976 | -28.5968
18.2823 | -35.098 | -28.6039
18.2884 | -35.0985 | -28.6112
18.2945 | -35.099 | -28.6184
18.3006 | -35.0995 | -28.6258
18.3067 | -35.1 | -28.633
18.3128 | -35.1005 | -28.6405
18.3189 | -35.1009 | -28.6478
18.325 | -35.1014 | -28.6549
18.3311 | -35.1019 | -28.6622
18.3372 | -35.1024 | -28.6696
18.3433 | -35.1029 | -28.6769
18.3494 | -35.1034 | -28.6841
18.3555 | -35.1039 | -28.6916
18.3616 | -35.1044 | -28.6989
18.3677 | -35.1048 | -28.7062
18.3738 | -35.1053 | -28.7136
18.3799 | -35.1058 | -28.721
18.386 | -35.1063 | -28.7283
18.3921 | -35.1068 | -28.7358
18.3982 | -35.1072 | -28.7428
18.4043 | -35.1077 | -28.75
18.4105 | -35.1082 | -28.7573
18.4166 | -35.1087 | -28.7647
18.4227 | -35.1091 | -28.7719
18.4288 | -35.1096 | -28.7794
18.4349 | -35.1101 | -28.7867
18.441 | -35.1106 | -28.794
18.4471 | -35.111 | -28.8011
18.4532 | -35.1115 | -28.8083
18.4593 | -35.112 | -28.8155
18.4654 | -35.1125 | -28.8226
18.4715 | -35.113 | -28.8301
18.4776 | -35.1134 | -28.8374
18.4837 | -35.1139 | -28.8448
18.4898 | -35.1144 | -28.8519
18.4959 | -35.1148 | -28.8593
18.502 | -35.1153 | -28.8666
18.5081 | -35.1158 | -28.8737
18.5142 | -35.1162 | -28.8811
18.5203 | -35.1168 | -28.8884
18.5264 | -35.1172 | -28.8956
18.5325 | -35.1177 | -28.9031
18.5386 | -35.1182 | -28.9105
18.5447 | -35.1187 | -28.9176
18.5508 | -35.1192 | -28.9249
18.557 | -35.1196 | -28.9324
18.5631 | -35.1201 | -28.9397
18.5692 | -35.1205 | -28.947
18.5753 | -35.121 | -28.9543
18.5814 | -35.1215 | -28.9615
18.5875 | -35.122 | -28.9689
18.5936 | -35.1225 | -28.976
18.5997 | -35.1229 | -28.9834
18.6058 | -35.1234 | -28.9903
18.6119 | -35.1238 | -28.9976
18.618 | -35.1243 | -29.0048
18.6241 | -35.1248 | -29.0122
18.6302 | -35.1253 | -29.0196
18.6363 | -35.1258 | -29.0269
18.6424 | -35.1263 | -29.0338
18.6485 | -35.1267 | -29.0411
18.6546 | -35.1272 | -29.0484
18.6607 | -35.1276 | -29.0558
18.6668 | -35.1281 | -29.0631
18.6729 | -35.1286 | -29.0703
18.679 | -35.1291 | -29.0772
18.6851 | -35.1295 | -29.0846
18.6912 | -35.13 | -29.0916
18.6974 | -35.1305 | -29.0989
18.7035 | -35.131 | -29.1061
18.7096 | -35.1314 | -29.1134
18.7157 | -35.1319 | -29.1205
18.7218 | -35.1324 | -29.1279
18.7279 | -35.1328 | -29.1351
18.734 | -35.1333 | -29.1425
18.7401 | -35.1338 | -29.1496
18.7462 | -35.1343 | -29.1565
18.7523 | -35.1348 | -29.1639
18.7584 | -35.1353 | -29.1713
18.7645 | -35.1357 | -29.1784
18.7706 | -35.1362 | -29.1856
18.7767 | -35.1367 | -29.1927
18.7828 | -35.1371 | -29.1997
18.7889 | -35.1376 | -29.2071
18.795 | -35.1381 | -29.2145
18.8011 | -35.1386 | -29.2217
18.8072 | -35.1391 | -29.2291
18.8133 | -35.1395 | -29.2365
18.8194 | -35.14 | -29.2437
18.8255 | -35.1405 | -29.2507
18.8316 | -35.1409 | -29.2578
18.8377 | -35.1414 | -29.2649
18.8439 | -35.1419 | -29.2724
18.85 | -35.1424 | -29.2796
18.8561 | -35.1428 | -29.2869
18.8622 | -35.1433 | -29.294
18.8683 | -35.1438 | -29.3012
18.8744 | -35.1442 | -29.3084
18.8805 | -35.1447 | -29.3159
18.8866 | -35.1452 | -29.3232
18.8927 | -35.1457 | -29.3302
18.8988 | -35.1462 | -29.3374
18.9049 | -35.1467 | -29.3445
18.911 | -35.1471 | -29.3518
18.9171 | -35.1476 | -29.3589
18.9232 | -35.1481 | -29.3659
18.9293 | -35.1486 | -29.373
18.9354 | -35.1491 | -29.3803
18.9415 | -35.1496 | -29.3874
18.9476 | -35.15 | -29.3947
18.9537 | -35.1505 | -29.4019
18.9598 | -35.1509 | -29.4087
18.9659 | -35.1514 | -29.4158
18.972 | -35.1518 | -29.423
18.9781 | -35.1523 | -29.4303
18.9843 | -35.1528 | -29.4376
18.9904 | -35.1533 | -29.4446
18.9965 | -35.1538 | -29.4517
19.0026 | -35.1543 | -29.4588
19.0087 | -35.1547 | -29.4661
19.0148 | -35.1552 | -29.4732
19.0209 | -35.1557 | -29.4804
19.027 | -35.1561 | -29.4876
19.0331 | -35.1566 | -29.4945
19.0392 | -35.1571 | -29.5018
19.0453 | -35.1576 | -29.5089
19.0514 | -35.1581 | -29.516
19.0575 | -35.1585 | -29.5233
19.0636 | -35.159 | -29.5301
19.0697 | -35.1595 | -29.5372
19.0758 | -35.1599 | -29.5443
19.0819 | -35.1604 | -29.5514
19.088 | -35.1609 | -29.5583
19.0941 | -35.1614 | -29.5657
19.1002 | -35.1618 | -29.5726
19.1063 | -35.1623 | -29.5797
19.1124 | -35.1627 | -29.5869
19.1185 | -35.1632 | -29.594
19.1246 | -35.1636 | -29.601
19.1308 | -35.1642 | -29.6081
19.1369 | -35.1647 | -29.615
19.143 | -35.1651 | -29.622
19.1491 | -35.1656 | -29.6291
19.1552 | -35.1661 | -29.6362
19.1613 | -35.1666 | -29.6436
19.1674 | -35.167 | -29.6509
19.1735 | -35.1675 | -29.6577
19.1796 | -35.168 | -29.6648
19.1857 | -35.1684 | -29.6723
19.1918 | -35.1689 | -29.6795
19.1979 | -35.1694 | -29.6867
19.204 | -35.1698 | -29.6938
19.2101 | -35.1704 | -29.7009
19.2162 | -35.1709 | -29.7079
19.2223 | -35.1713 | -29.7151
19.2284 | -35.1718 | -29.7222
19.2345 | -35.1723 | -29.7293
19.2406 | -35.1728 | -29.7364
19.2467 | -35.1732 | -29.7438
19.2528 | -35.1737 | -29.7509
19.2589 | -35.1742 | -29.7579
19.265 | -35.1746 | -29.765
19.2712 | -35.1751 | -29.7721
19.2773 | -35.1755 | -29.7793
19.2834 | -35.176 | -29.7864
19.2895 | -35.1765 | -29.7936
19.2956 | -35.177 | -29.8008
19.3017 | -35.1774 | -29.8079
19.3078 | -35.1778 | -29.8152
19.3139 | -35.1783 | -29.8223
19.32 | -35.1787 | -29.8293
19.3261 | -35.1792 | -29.8363
19.3322 | -35.1797 | -29.8434
19.3383 | -35.1802 | -29.8506
19.3444 | -35.1807 | -29.8577
19.3505 | -35.1811 | -29.865
19.3566 | -35.1815 | -29.8721
19.3627 | -35.182 | -29.8793
19.3688 | -35.1825 | -29.8864
19.3749 | -35.183 | -29.8934
19.381 | -35.1835 | -29.9
19.3871 | -35.184 | -29.9071
19.3932 | -35.1844 | -29.9144
19.3993 | -35.1849 | -29.9213
19.4054 | -35.1853 | -29.9285
19.4115 | -35.1858 | -29.9357
19.4177 | -35.1862 | -29.9426
19.4238 | -35.1867 | -29.9496
19.4299 | -35.1871 | -29.9567
19.436 | -35.1877 | -29.964
19.4421 | -35.1881 | -29.9709
19.4482 | -35.1886 | -29.9779
19.4543 | -35.1891 | -29.9846
19.4604 | -35.1895 | -29.9916
19.4665 | -35.19 | -29.9989
19.4726 | -35.1904 | -30.0057
19.4787 | -35.1909 | -30.0125
19.4848 | -35.1914 | -30.0195
19.4909 | -35.1919 | -30.0262
19.497 | -35.1923 | -30.0335
19.5031 | -35.1928 | -30.0402
19.5092 | -35.1932 | -30.0475
19.5153 | -35.1937 | -30.0544
19.5214 | -35.1942 | -30.0611
19.5275 | -35.1946 | -30.0684
19.5336 | -35.1951 | -30.0754
19.5397 | -35.1956 | -30.0824
19.5458 | -35.196 | -30.0894
19.5519 | -35.1965 | -30.0965
19.5581 | -35.197 | -30.1037
19.5642 | -35.1975 | -30.1107
19.5703 | -35.1979 | -30.1176
19.5764 | -35.1983 | -30.1245
19.5825 | -35.1988 | -30.1314
19.5886 | -35.1992 | -30.1385
19.5947 | -35.1997 | -30.145
19.6008 | -35.2002 | -30.1518
19.6069 | -35.2006 | -30.1586
19.613 | -35.2011 | -30.1658
19.6191 | -35.2016 | -30.1729
19.6252 | -35.202 | -30.1796
19.6313 | -35.2025 | -30.1862
19.6374 | -35.2029 | -30.1934
19.6435 | -35.2034 | -30.2006
19.6496 | -35.2039 | -30.2073
19.6557 | -35.2044 | -30.2143
19.6618 | -35.2048 | -30.2211
19.6679 | -35.2053 | -30.2279
19.674 | -35.2057 | -30.2347
19.6801 | -35.2062 | -30.2415
19.6862 | -35.2067 | -30.2485
19.6923 | -35.2071 | -30.2553
19.6984 | -35.2076 | -30.262
19.7046 | -35.2081 | -30.2689
19.7107 | -35.2085 | -30.2758
19.7168 | -35.209 | -30.2828
19.7229 | -35.2095 | -30.2897
19.729 | -35.2099 | -30.2969
19.7351 | -35.2104 | -30.3037
19.7412 | -35.2109 | -30.3108
19.7473 | -35.2114 | -30.3177
19.7534 | -35.2118 | -30.3247
19.7595 | -35.2123 | -30.3316
19.7656 | -35.2128 | -30.3389
19.7717 | -35.2132 | -30.3454
19.7778 | -35.2137 | -30.3524
19.7839 | -35.2142 | -30.3593
19.79 | -35.2146 | -30.3666
19.7961 | -35.2151 | -30.3736
19.8022 | -35.2155 | -30.3804
19.8083 | -35.216 | -30.3873
19.8144 | -35.2164 | -30.3942
19.8205 | -35.2168 | -30.4013
19.8266 | -35.2173 | -30.408
19.8327 | -35.2178 | -30.4149
19.8388 | -35.2183 | -30.4214
19.845 | -35.2187 | -30.4282
19.8511 | -35.2192 | -30.4353
19.8572 | -35.2196 | -30.4421
19.8633 | -35.2201 | -30.449
19.8694 | -35.2206 | -30.4558
19.8755 | -35.221 | -30.4628
19.8816 | -35.2215 | -30.4698
19.8877 | -35.222 | -30.4767
19.8938 | -35.2224 | -30.4837
19.8999 | -35.2229 | -30.4905
19.906 | -35.2234 | -30.4974
19.9121 | -35.2239 | -30.5041
19.9182 | -35.2244 | -30.511
19.9243 | -35.2248 | -30.5179
19.9304 | -35.2253 | -30.5247
19.9365 | -35.2258 | -30.5319
19.9426 | -35.2263 | -30.5383
19.9487 | -35.2268 | -30.5452
19.9548 | -35.2272 | -30.5523
19.9609 | -35.2277 | -30.5594
19.967 | -35.2281 | -30.5661
19.9731 | -35.2286 | -30.573
19.9792 | -35.229 | -30.5796
19.9853 | -35.2294 | -30.5864
19.9915 | -35.2299 | -30.5934
19.9976 | -35.2303 | -30.6003
20.0037 | -35.2308 | -30.6072
20.0098 | -35.2313 | -30.6141
20.0159 | -35.2317 | -30.6205
20.022 | -35.2321 | -30.6275
20.0281 | -35.2326 | -30.6346
20.0342 | -35.2331 | -30.6413
20.0403 | -35.2335 | -30.6484
20.0464 | -35.2339 | -30.6551
20.0525 | -35.2344 | -30.6621
20.0586 | -35.2349 | -30.6691
20.0647 | -35.2354 | -30.6762
20.0708 | -35.2358 | -30.6833
20.0769 | -35.2363 | -30.6902
20.083 | -35.2368 | -30.6964
20.0891 | -35.2372 | -30.7031
20.0952 | -35.2377 | -30.7097
20.1013 | -35.2382 | -30.7163
20.1074 | -35.2387 | -30.7232
20.1135 | -35.2391 | -30.7301
20.1196 | -35.2396 | -30.7368
20.1257 | -35.24 | -30.7437
20.1319 | -35.2405 | -30.7504
20.138 | -35.2409 | -30.7572
20.1441 | -35.2414 | -30.7641
20.1502 | -35.2419 | -30.7708
20.1563 | -35.2423 | -30.7774
20.1624 | -35.2427 | -30.7842
20.1685 | -35.2432 | -30.791
20.1746 | -35.2436 | -30.7977
20.1807 | -35.244 | -30.8047
20.1868 | -35.2445 | -30.8114
20.1929 | -35.2449 | -30.818
20.199 | -35.2454 | -30.8249
20.2051 | -35.2458 | -30.8316
20.2112 | -35.2462 | -30.8386
20.2173 | -35.2466 | -30.8452
20.2234 | -35.2471 | -30.8522
20.2295 | -35.2475 | -30.8589
20.2356 | -35.2479 | -30.8658
20.2417 | -35.2484 | -30.8725
20.2478 | -35.2488 | -30.879
20.2539 | -35.2493 | -30.886
20.26 | -35.2498 | -30.892
20.2661 | -35.2502 | -30.8989
20.2723 | -35.2507 | -30.9059
20.2784 | -35.2511 | -30.9127
20.2845 | -35.2515 | -30.9194
20.2906 | -35.252 | -30.9261
20.2967 | -35.2524 | -30.9328
20.3028 | -35.2528 | -30.9397
20.3089 | -35.2533 | -30.9463
20.315 | -35.2538 | -30.953
20.3211 | -35.2542 | -30.9598
20.3272 | -35.2547 | -30.9665
20.3333 | -35.2551 | -30.9732
20.3394 | -35.2556 | -30.9798
20.3455 | -35.2561 | -30.9865
20.3516 | -35.2566 | -30.9931
20.3577 | -35.257 | -30.9998
20.3638 | -35.2574 | -31.0066
20.3699 | -35.2579 | -31.0131
20.376 | -35.2584 | -31.02
20.3821 | -35.2588 | -31.0267
20.3882 | -35.2593 | -31.0334
20.3943 | -35.2598 | -31.0402
20.4004 | -35.2602 | -31.0467
20.4065 | -35.2607 | -31.0535
20.4126 | -35.2611 | -31.0603
20.4188 | -35.2616 | -31.067
20.4249 | -35.2621 | -31.0737
20.431 | -35.2625 | -31.0804
20.4371 | -35.263 | -31.0872
20.4432 | -35.2634 | -31.0939
20.4493 | -35.2639 | -31.1007
20.4554 | -35.2644 | -31.1072
20.4615 | -35.2649 | -31.1141
20.4676 | -35.2653 | -31.1211
20.4737 | -35.2658 | -31.128
20.4798 | -35.2662 | -31.1346
20.4859 | -35.2666 | -31.1416
20.492 | -35.267 | -31.1479
20.4981 | -35.2675 | -31.1544
20.5042 | -35.2679 | -31.1612
20.5103 | -35.2684 | -31.1678
20.5164 | -35.2688 | -31.1745
20.5225 | -35.2692 | -31.1813
20.5286 | -35.2697 | -31.188
20.5347 | -35.2701 | -31.1946
20.5408 | -35.2705 | -31.2009
20.5469 | -35.271 | -31.2076
20.553 | -35.2715 | -31.2145
20.5592 | -35.2719 | -31.2213
20.5653 | -35.2723 | -31.2281
20.5714 | -35.2728 | -31.2348
20.5775 | -35.2733 | -31.2414
20.5836 | -35.2737 | -31.2483
20.5897 | -35.2741 | -31.2551
20.5958 | -35.2746 | -31.2619
20.6019 | -35.275 | -31.2688
20.608 | -35.2755 | -31.2755
20.6141 | -35.276 | -31.282
20.6202 | -35.2764 | -31.2884
20.6263 | -35.2768 | -31.2951
20.6324 | -35.2773 | -31.3016
20.6385 | -35.2778 | -31.3083
20.6446 | -35.2782 | -31.3153
20.6507 | -35.2787 | -31.322
20.6568 | -35.2791 | -31.3285
20.6629 | -35.2795 | -31.3356
20.669 | -35.28 | -31.3422
20.6751 | -35.2805 | -31.3488
20.6812 | -35.281 | -31.3553
20.6873 | -35.2815 | -31.3619
20.6934 | -35.2819 | -31.3683
20.6995 | -35.2823 | -31.3747
20.7057 | -35.2828 | -31.3813
20.7118 | -35.2832 | -31.3882
20.7179 | -35.2836 | -31.3947
20.724 | -35.284 | -31.4014
20.7301 | -35.2845 | -31.4081
20.7362 | -35.2849 | -31.415
20.7423 | -35.2854 | -31.4219
20.7484 | -35.2858 | -31.4286
20.7545 | -35.2862 | -31.4349
20.7606 | -35.2867 | -31.4414
20.7667 | -35.2872 | -31.4482
20.7728 | -35.2877 | -31.4547
20.7789 | -35.2881 | -31.4614
20.785 | -35.2886 | -31.4678
20.7911 | -35.2891 | -31.4745
20.7972 | -35.2895 | -31.4808
20.8033 | -35.29 | -31.4874
20.8094 | -35.2904 | -31.4943
20.8155 | -35.2909 | -31.5005
20.8216 | -35.2913 | -31.507
20.8277 | -35.2917 | -31.5132
20.8338 | -35.2921 | -31.5198
20.8399 | -35.2925 | -31.5265
20.8461 | -35.293 | -31.5329
20.8522 | -35.2934 | -31.5398
20.8583 | -35.2938 | -31.5462
20.8644 | -35.2943 | -31.553
20.8705 | -35.2947 | -31.5596
20.8766 | -35.2952 | -31.5664
20.8827 | -35.2956 | -31.573
20.8888 | -35.2961 | -31.5794
20.8949 | -35.2964 | -31.586
20.901 | -35.2969 | -31.5926
20.9071 | -35.2973 | -31.599
20.9132 | -35.2978 | -31.6059
20.9193 | -35.2982 | -31.6127
20.9254 | -35.2987 | -31.6197
20.9315 | -35.2991 | -31.6261
20.9376 | -35.2995 | -31.6326
20.9437 | -35.2999 | -31.6392
20.9498 | -35.3004 | -31.646
20.9559 | -35.3008 | -31.6528
20.962 | -35.3012 | -31.6593
20.9681 | -35.3017 | -31.6658
20.9742 | -35.3022 | -31.6724
20.9803 | -35.3027 | -31.679
20.9864 | -35.3031 | -31.6854
20.9926 | -35.3035 | -31.6922
20.9987 | -35.3039 | -31.6987
21.0048 | -35.3044 | -31.7054
21.0109 | -35.3049 | -31.7121
21.017 | -35.3053 | -31.7187
21.0231 | -35.3057 | -31.7253
21.0292 | -35.3062 | -31.7319
21.0353 | -35.3066 | -31.7384
21.0414 | -35.307 | -31.7447
21.0475 | -35.3074 | -31.7513
21.0536 | -35.3079 | -31.7577
21.0597 | -35.3083 | -31.7643
21.0658 | -35.3088 | -31.771
21.0719 | -35.3092 | -31.7774
21.078 | -35.3096 | -31.7842
21.0841 | -35.3101 | -31.7908
21.0902 | -35.3105 | -31.7972
21.0963 | -35.311 | -31.8039
21.1024 | -35.3115 | -31.8103
21.1085 | -35.3119 | -31.8171
21.1146 | -35.3124 | -31.8236
21.1207 | -35.3128 | -31.8305
21.1268 | -35.3133 | -31.8372
21.133 | -35.3138 | -31.844
21.1391 | -35.3143 | -31.8507
21.1452 | -35.3147 | -31.8571
21.1513 | -35.3151 | -31.864
21.1574 | -35.3156 | -31.8708
21.1635 | -35.316 | -31.8771
21.1696 | -35.3164 | -31.8832
21.1757 | -35.3168 | -31.8892
21.1818 | -35.3172 | -31.8959
21.1879 | -35.3177 | -31.9026
21.194 | -35.3181 | -31.9092
21.2001 | -35.3186 | -31.9154
21.2062 | -35.319 | -31.922
21.2123 | -35.3195 | -31.9286
21.2184 | -35.3199 | -31.9349
21.2245 | -35.3203 | -31.9415
21.2306 | -35.3208 | -31.948
21.2367 | -35.3212 | -31.9544
21.2428 | -35.3217 | -31.9609
21.2489 | -35.3221 | -31.9672
21.255 | -35.3225 | -31.9738
21.2611 | -35.3229 | -31.9803
21.2672 | -35.3233 | -31.9867
21.2733 | -35.3237 | -31.9931
21.2795 | -35.3241 | -31.9993
21.2856 | -35.3246 | -32.0056
21.2917 | -35.3251 | -32.0121
21.2978 | -35.3255 | -32.0186
21.3039 | -35.326 | -32.0251
21.31 | -35.3265 | -32.0318
21.3161 | -35.3269 | -32.0382
21.3222 | -35.3273 | -32.0453
21.3283 | -35.3277 | -32.0517
21.3344 | -35.3282 | -32.0585
21.3405 | -35.3286 | -32.0651
21.3466 | -35.3291 | -32.0713
21.3527 | -35.3296 | -32.0779
21.3588 | -35.33 | -32.0843
21.3649 | -35.3304 | -32.0909
21.371 | -35.3309 | -32.0974
21.3771 | -35.3312 | -32.1036
21.3832 | -35.3316 | -32.11
21.3893 | -35.3321 | -32.1163
21.3954 | -35.3325 | -32.1226
21.4015 | -35.3329 | -32.1289
21.4076 | -35.3334 | -32.1355
21.4137 | -35.3338 | -32.1421
21.4199 | -35.3343 | -32.1485
21.426 | -35.3347 | -32.1548
21.4321 | -35.3351 | -32.1609
21.4382 | -35.3356 | -32.1674
21.4443 | -35.336 | -32.1738
21.4504 | -35.3364 | -32.1803
21.4565 | -35.3368 | -32.1863
21.4626 | -35.3372 | -32.1928
21.4687 | -35.3376 | -32.1994
21.4748 | -35.3381 | -32.2057
21.4809 | -35.3385 | -32.2125
21.487 | -35.3389 | -32.2188
21.4931 | -35.3394 | -32.2254
21.4992 | -35.3398 | -32.2317
21.5053 | -35.3403 | -32.2381
21.5114 | -35.3406 | -32.2443
21.5175 | -35.3411 | -32.2508
21.5236 | -35.3415 | -32.2576
21.5297 | -35.3419 | -32.2639
21.5358 | -35.3423 | -32.27
21.5419 | -35.3428 | -32.2769
21.548 | -35.3432 | -32.2831
21.5541 | -35.3437 | -32.2898
21.5602 | -35.3441 | -32.2962
21.5664 | -35.3446 | -32.3027
21.5725 | -35.345 | -32.3092
21.5786 | -35.3455 | -32.3155
21.5847 | -35.3459 | -32.3223
21.5908 | -35.3463 | -32.3288
21.5969 | -35.3467 | -32.3348
21.603 | -35.3471 | -32.3417
21.6091 | -35.3476 | -32.3481
21.6152 | -35.348 | -32.3548
21.6213 | -35.3484 | -32.3612
21.6274 | -35.3488 | -32.3675
21.6335 | -35.3492 | -32.3739
21.6396 | -35.3496 | -32.3805
21.6457 | -35.3501 | -32.3873
21.6518 | -35.3505 | -32.3938
21.6579 | -35.3509 | -32.4002
21.664 | -35.3513 | -32.4064
21.6701 | -35.3518 | -32.4128
21.6762 | -35.3522 | -32.4192
21.6823 | -35.3526 | -32.4254
21.6884 | -35.3531 | -32.432
21.6945 | -35.3536 | -32.4384
21.7006 | -35.354 | -32.4448
21.7068 | -35.3544 | -32.4512
21.7129 | -35.3549 | -32.4579
21.719 | -35.3553 | -32.4645
21.7251 | -35.3557 | -32.4707
21.7312 | -35.3562 | -32.4769
21.7373 | -35.3566 | -32.4833
21.7434 | -35.357 | -32.4897
21.7495 | -35.3575 | -32.4964
21.7556 | -35.3579 | -32.503
21.7617 | -35.3584 | -32.5094
21.7678 | -35.3588 | -32.5159
21.7739 | -35.3592 | -32.5218
21.78 | -35.3597 | -32.5282
21.7861 | -35.3601 | -32.5347
21.7922 | -35.3606 | -32.5411
21.7983 | -35.3611 | -32.5476
21.8044 | -35.3615 | -32.5536
21.8105 | -35.362 | -32.5595
21.8166 | -35.3624 | -32.5661
21.8227 | -35.3629 | -32.5726
21.8288 | -35.3633 | -32.5786
21.8349 | -35.3638 | -32.5855
21.841 | -35.3642 | -32.5922
21.8471 | -35.3647 | -32.5987
21.8533 | -35.3652 | -32.605
21.8594 | -35.3657 | -32.6117
21.8655 | -35.3661 | -32.618
21.8716 | -35.3666 | -32.6245
21.8777 | -35.367 | -32.6308
21.8838 | -35.3675 | -32.6371
21.8899 | -35.3679 | -32.6437
21.896 | -35.3684 | -32.6503
21.9021 | -35.3688 | -32.6565
21.9082 | -35.3692 | -32.6629
21.9143 | -35.3697 | -32.6691
21.9204 | -35.3701 | -32.6756
21.9265 | -35.3706 | -32.6818
21.9326 | -35.371 | -32.6883
21.9387 | -35.3715 | -32.6947
21.9448 | -35.3719 | -32.701
21.9509 | -35.3723 | -32.7072
21.957 | -35.3728 | -32.7137
21.9631 | -35.3732 | -32.7202
21.9692 | -35.3736 | -32.7265
21.9753 | -35.3741 | -32.733
21.9814 | -35.3745 | -32.7396
21.9875 | -35.3749 | -32.7457
21.9937 | -35.3753 | -32.7522
21.9998 | -35.3757 | -32.7589
22.0059 | -35.3761 | -32.7651
22.012 | -35.3766 | -32.7715
22.0181 | -35.3771 | -32.7779
22.0242 | -35.3775 | -32.7843
22.0303 | -35.3779 | -32.7911
22.0364 | -35.3783 | -32.7977
22.0425 | -35.3788 | -32.804
22.0486 | -35.3792 | -32.8107
22.0547 | -35.3797 | -32.8174
22.0608 | -35.38 | -32.8235
22.0669 | -35.3804 | -32.8302
22.073 | -35.3809 | -32.8366
22.0791 | -35.3813 | -32.8427
22.0852 | -35.3818 | -32.8492
22.0913 | -35.3822 | -32.8556
22.0974 | -35.3827 | -32.8621
22.1035 | -35.3831 | -32.8683
22.1096 | -35.3835 | -32.8747
22.1157 | -35.384 | -32.8809
22.1218 | -35.3844 | -32.8875
22.1279 | -35.3849 | -32.8937
22.134 | -35.3853 | -32.9001
22.1402 | -35.3857 | -32.9064
22.1463 | -35.3861 | -32.9128
22.1524 | -35.3866 | -32.9193
22.1585 | -35.387 | -32.9259
22.1646 | -35.3874 | -32.9318
22.1707 | -35.3879 | -32.9383
22.1768 | -35.3884 | -32.9445
22.1829 | -35.3888 | -32.9509
22.189 | -35.3892 | -32.9574
22.1951 | -35.3896 | -32.9639
22.2012 | -35.3901 | -32.9702
22.2073 | -35.3905 | -32.9765
22.2134 | -35.3909 | -32.9827
22.2195 | -35.3913 | -32.9891
22.2256 | -35.3918 | -32.9954
22.2317 | -35.3922 | -33.0017
22.2378 | -35.3926 | -33.0082
22.2439 | -35.393 | -33.0145
22.25 | -35.3935 | -33.021
22.2561 | -35.3939 | -33.0278
22.2622 | -35.3943 | -33.0341
22.2683 | -35.3948 | -33.0402
22.2744 | -35.3952 | -33.0468
22.2806 | -35.3956 | -33.0535
22.2867 | -35.3961 | -33.0599
22.2928 | -35.3965 | -33.0662
22.2989 | -35.3969 | -33.0724
22.305 | -35.3974 | -33.0789
22.3111 | -35.3979 | -33.0856
22.3172 | -35.3983 | -33.0918
22.3233 | -35.3988 | -33.0984
22.3294 | -35.3992 | -33.1047
22.3355 | -35.3996 | -33.1111
22.3416 | -35.4001 | -33.1178
22.3477 | -35.4005 | -33.1239
22.3538 | -35.4009 | -33.1303
22.3599 | -35.4013 | -33.1367
22.366 | -35.4017 | -33.1427
22.3721 | -35.4021 | -33.1488
22.3782 | -35.4025 | -33.1549
22.3843 | -35.4029 | -33.1611
22.3904 | -35.4033 | -33.1673
22.3965 | -35.4037 | -33.1739
22.4026 | -35.4042 | -33.1802
22.4087 | -35.4046 | -33.1869
22.4148 | -35.405 | -33.1933
22.4209 | -35.4054 | -33.2
22.4271 | -35.4059 | -33.2062
22.4332 | -35.4063 | -33.2125
22.4393 | -35.4067 | -33.2194
22.4454 | -35.4071 | -33.2258
22.4515 | -35.4076 | -33.2325
22.4576 | -35.408 | -33.239
22.4637 | -35.4084 | -33.2456
22.4698 | -35.4089 | -33.2525
22.4759 | -35.4093 | -33.2586
22.482 | -35.4098 | -33.2652
22.4881 | -35.4102 | -33.2716
22.4942 | -35.4106 | -33.2782
22.5003 | -35.411 | -33.2843
22.5064 | -35.4114 | -33.2908
22.5125 | -35.4118 | -33.2971
22.5186 | -35.4123 | -33.3034
22.5247 | -35.4127 | -33.31
22.5308 | -35.4131 | -33.3163
22.5369 | -35.4135 | -33.3229
22.543 | -35.414 | -33.3292
22.5491 | -35.4144 | -33.3348
22.5552 | -35.4148 | -33.3413
22.5613 | -35.4152 | -33.3476
22.5675 | -35.4156 | -33.3539
22.5736 | -35.4161 | -33.3602
22.5797 | -35.4165 | -33.367
22.5858 | -35.4169 | -33.3731
22.5919 | -35.4173 | -33.3796
22.598 | -35.4178 | -33.3858
22.6041 | -35.4182 | -33.3926
22.6102 | -35.4187 | -33.3989
22.6163 | -35.4191 | -33.4049
22.6224 | -35.4196 | -33.4116
22.6285 | -35.42 | -33.4183
22.6346 | -35.4204 | -33.4249
22.6407 | -35.4209 | -33.4312
22.6468 | -35.4214 | -33.4376
22.6529 | -35.4218 | -33.4439
22.659 | -35.4223 | -33.4503
22.6651 | -35.4227 | -33.4564
22.6712 | -35.423 | -33.4626
22.6773 | -35.4234 | -33.4691
22.6834 | -35.4239 | -33.4759
22.6895 | -35.4243 | -33.4821
22.6956 | -35.4248 | -33.4886
22.7017 | -35.4252 | -33.4948
22.7079 | -35.4256 | -33.5008
22.714 | -35.426 | -33.5073
22.7201 | -35.4265 | -33.5137
22.7262 | -35.427 | -33.5202
22.7323 | -35.4274 | -33.5263
22.7384 | -35.4277 | -33.5327
22.7445 | -35.4282 | -33.5391
22.7506 | -35.4286 | -33.5451
22.7567 | -35.4291 | -33.5515
22.7628 | -35.4294 | -33.5581
22.7689 | -35.4298 | -33.5643
22.775 | -35.4303 | -33.5705
22.7811 | -35.4307 | -33.5766
22.7872 | -35.4311 | -33.5832
22.7933 | -35.4315 | -33.5892
22.7994 | -35.4319 | -33.5958
22.8055 | -35.4324 | -33.6019
22.8116 | -35.4329 | -33.6086
22.8177 | -35.4333 | -33.6152
22.8238 | -35.4338 | -33.6214
22.8299 | -35.4343 | -33.6275
22.836 | -35.4346 | -33.6335
22.8421 | -35.435 | -33.6396
22.8482 | -35.4354 | -33.6458
22.8544 | -35.4358 | -33.652
22.8605 | -35.4362 | -33.658
22.8666 | -35.4366 | -33.6647
22.8727 | -35.4371 | -33.6712
22.8788 | -35.4375 | -33.6775
22.8849 | -35.4379 | -33.6841
22.891 | -35.4383 | -33.6903
22.8971 | -35.4388 | -33.6968
22.9032 | -35.4392 | -33.7026
22.9093 | -35.4397 | -33.7089
22.9154 | -35.4401 | -33.7152
22.9215 | -35.4405 | -33.7215
22.9276 | -35.4409 | -33.7279
22.9337 | -35.4413 | -33.7343
22.9398 | -35.4417 | -33.7407
22.9459 | -35.4422 | -33.7472
22.952 | -35.4426 | -33.7534
22.9581 | -35.443 | -33.76
22.9642 | -35.4435 | -33.7662
22.9703 | -35.4439 | -33.7726
22.9764 | -35.4444 | -33.779
22.9825 | -35.4448 | -33.7852
22.9886 | -35.4453 | -33.791
22.9948 | -35.4457 | -33.7972
23.0009 | -35.4461 | -33.8034
23.007 | -35.4465 | -33.8094
23.0131 | -35.4469 | -33.8157
23.0192 | -35.4473 | -33.8216
23.0253 | -35.4478 | -33.8277
23.0314 | -35.4482 | -33.8337
23.0375 | -35.4486 | -33.8398
23.0436 | -35.449 | -33.8462
23.0497 | -35.4494 | -33.8527
23.0558 | -35.4498 | -33.8591
23.0619 | -35.4503 | -33.8657
23.068 | -35.4507 | -33.8723
23.0741 | -35.4511 | -33.8787
23.0802 | -35.4515 | -33.8853
23.0863 | -35.452 | -33.8916
23.0924 | -35.4524 | -33.8976
23.0985 | -35.4529 | -33.9039
23.1046 | -35.4533 | -33.9105
23.1107 | -35.4537 | -33.9166
23.1168 | -35.4542 | -33.9226
23.1229 | -35.4546 | -33.9286
23.129 | -35.455 | -33.9348
23.1351 | -35.4555 | -33.9409
23.1413 | -35.4559 | -33.9472
23.1474 | -35.4563 | -33.9535
23.1535 | -35.4567 | -33.9597
23.1596 | -35.4572 | -33.9656
23.1657 | -35.4576 | -33.9721
23.1718 | -35.4581 | -33.9783
23.1779 | -35.4585 | -33.9844
23.184 | -35.4589 | -33.9907
23.1901 | -35.4594 | -33.9969
23.1962 | -35.4598 | -34.0031
23.2023 | -35.4602 | -34.0097
23.2084 | -35.4607 | -34.0159
23.2145 | -35.461 | -34.022
23.2206 | -35.4614 | -34.0281
23.2267 | -35.4619 | -34.0345
23.2328 | -35.4623 | -34.0407
23.2389 | -35.4628 | -34.047
23.245 | -35.4632 | -34.0533
23.2511 | -35.4636 | -34.0596
23.2572 | -35.4641 | -34.0659
23.2633 | -35.4645 | -34.072
23.2694 | -35.4649 | -34.078
23.2755 | -35.4653 | -34.0844
23.2817 | -35.4657 | -34.0907
23.2878 | -35.4661 | -34.0968
23.2939 | -35.4665 | -34.1032
23.3 | -35.467 | -34.1095
23.3061 | -35.4674 | -34.1159
23.3122 | -35.4679 | -34.1222
23.3183 | -35.4683 | -34.1286
23.3244 | -35.4687 | -34.1349
23.3305 | -35.4691 | -34.1414
23.3366 | -35.4695 | -34.1478
23.3427 | -35.4699 | -34.1539
23.3488 | -35.4704 | -34.1602
23.3549 | -35.4708 | -34.1664
23.361 | -35.4712 | -34.1723
23.3671 | -35.4716 | -34.1782
23.3732 | -35.4721 | -34.1844
23.3793 | -35.4725 | -34.1905
23.3854 | -35.4729 | -34.1967
23.3915 | -35.4733 | -34.2032
23.3976 | -35.4737 | -34.2092
23.4037 | -35.4741 | -34.2156
23.4098 | -35.4746 | -34.2216
23.4159 | -35.475 | -34.228
23.422 | -35.4754 | -34.2343
23.4282 | -35.4759 | -34.2406
23.4343 | -35.4763 | -34.2469
23.4404 | -35.4767 | -34.2535
23.4465 | -35.4771 | -34.2598
23.4526 | -35.4776 | -34.2663
23.4587 | -35.478 | -34.2726
23.4648 | -35.4785 | -34.2786
23.4709 | -35.4788 | -34.2849
23.477 | -35.4793 | -34.2912
23.4831 | -35.4797 | -34.297
23.4892 | -35.4801 | -34.3031
23.4953 | -35.4806 | -34.3092
23.5014 | -35.481 | -34.3156
23.5075 | -35.4814 | -34.3219
23.5136 | -35.4819 | -34.3281
23.5197 | -35.4823 | -34.3341
23.5258 | -35.4827 | -34.3402
23.5319 | -35.4832 | -34.3461
23.538 | -35.4836 | -34.3523
23.5441 | -35.484 | -34.3585
23.5502 | -35.4845 | -34.3648
23.5563 | -35.4849 | -34.3709
23.5624 | -35.4853 | -34.3772
23.5686 | -35.4858 | -34.3829
23.5747 | -35.4862 | -34.3891
23.5808 | -35.4867 | -34.3952
23.5869 | -35.4871 | -34.4012
23.593 | -35.4875 | -34.4078
23.5991 | -35.4879 | -34.4141
23.6052 | -35.4883 | -34.4206
23.6113 | -35.4888 | -34.4269
23.6174 | -35.4892 | -34.433
23.6235 | -35.4896 | -34.4393
23.6296 | -35.4901 | -34.4457
23.6357 | -35.4905 | -34.4519
23.6418 | -35.4909 | -34.4579
23.6479 | -35.4913 | -34.4642
23.654 | -35.4917 | -34.4705
23.6601 | -35.4922 | -34.4768
23.6662 | -35.4926 | -34.4832
23.6723 | -35.493 | -34.4898
23.6784 | -35.4935 | -34.4961
23.6845 | -35.4939 | -34.5024
23.6906 | -35.4944 | -34.5084
23.6967 | -35.4948 | -34.5148
23.7028 | -35.4953 | -34.5208
23.7089 | -35.4957 | -34.5266
23.7151 | -35.4961 | -34.533
23.7212 | -35.4965 | -34.5393
23.7273 | -35.4968 | -34.5456
23.7334 | -35.4973 | -34.5518
23.7395 | -35.4977 | -34.5581
23.7456 | -35.4981 | -34.564
23.7517 | -35.4985 | -34.5705
23.7578 | -35.4989 | -34.5766
23.7639 | -35.4993 | -34.5829
23.77 | -35.4997 | -34.5891
23.7761 | -35.5001 | -34.5951
23.7822 | -35.5005 | -34.6013
23.7883 | -35.5009 | -34.6077
23.7944 | -35.5014 | -34.6139
23.8005 | -35.5018 | -34.62
23.8066 | -35.5022 | -34.6263
23.8127 | -35.5026 | -34.6322
23.8188 | -35.503 | -34.6384
23.8249 | -35.5034 | -34.6447
23.831 | -35.5038 | -34.651
23.8371 | -35.5042 | -34.6571
23.8432 | -35.5046 | -34.6634
23.8493 | -35.5051 | -34.6695
23.8555 | -35.5055 | -34.6756
23.8616 | -35.5059 | -34.6818
23.8677 | -35.5063 | -34.6878
23.8738 | -35.5067 | -34.6939
23.8799 | -35.5071 | -34.7005
23.886 | -35.5076 | -34.7064
23.8921 | -35.508 | -34.7125
23.8982 | -35.5084 | -34.7186
23.9043 | -35.5088 | -34.7247
23.9104 | -35.5092 | -34.7309
23.9165 | -35.5096 | -34.7372
23.9226 | -35.51 | -34.7429
23.9287 | -35.5104 | -34.7492
23.9348 | -35.5109 | -34.7557
23.9409 | -35.5113 | -34.7619
23.947 | -35.5117 | -34.7679
23.9531 | -35.5121 | -34.7742
23.9592 | -35.5125 | -34.7803
23.9653 | -35.513 | -34.7863
23.9714 | -35.5134 | -34.7923
23.9775 | -35.5138 | -34.7987
23.9836 | -35.5142 | -34.8051
23.9897 | -35.5146 | -34.8112
23.9958 | -35.515 | -34.8175
24.002 | -35.5154 | -34.8236
24.0081 | -35.5158 | -34.8297
24.0142 | -35.5162 | -34.8358
24.0203 | -35.5167 | -34.8422
24.0264 | -35.5171 | -34.8486
24.0325 | -35.5175 | -34.8545
24.0386 | -35.5179 | -34.8606
24.0447 | -35.5183 | -34.8666
24.0508 | -35.5187 | -34.8729
24.0569 | -35.5191 | -34.8792
24.063 | -35.5195 | -34.8851
24.0691 | -35.5199 | -34.8913
24.0752 | -35.5203 | -34.8974
24.0813 | -35.5207 | -34.9035
24.0874 | -35.5211 | -34.9095
24.0935 | -35.5215 | -34.9154
24.0996 | -35.5219 | -34.9211
24.1057 | -35.5223 | -34.9272
24.1118 | -35.5228 | -34.9332
24.1179 | -35.5232 | -34.9394
24.124 | -35.5236 | -34.9457
24.1301 | -35.524 | -34.9518
24.1362 | -35.5244 | -34.958
24.1424 | -35.5248 | -34.964
24.1485 | -35.5252 | -34.9702
24.1546 | -35.5256 | -34.9759
24.1607 | -35.5259 | -34.982
24.1668 | -35.5263 | -34.9878
24.1729 | -35.5268 | -34.9942
24.179 | -35.5272 | -35.0004
24.1851 | -35.5275 | -35.0066
24.1912 | -35.528 | -35.0127
24.1973 | -35.5284 | -35.0188
24.2034 | -35.5288 | -35.0244
24.2095 | -35.5292 | -35.0303
24.2156 | -35.5296 | -35.0367
24.2217 | -35.53 | -35.0429
24.2278 | -35.5304 | -35.0488
24.2339 | -35.5308 | -35.055
24.24 | -35.5312 | -35.0611
24.2461 | -35.5317 | -35.0671
24.2522 | -35.532 | -35.0733
24.2583 | -35.5325 | -35.0792
24.2644 | -35.5329 | -35.0853
24.2705 | -35.5334 | -35.0913
24.2766 | -35.5337 | -35.0972
24.2827 | -35.5342 | -35.1033
24.2889 | -35.5346 | -35.1094
24.295 | -35.535 | -35.1157
24.3011 | -35.5354 | -35.1222
24.3072 | -35.5358 | -35.128
24.3133 | -35.5362 | -35.1339
24.3194 | -35.5366 | -35.1399
24.3255 | -35.5371 | -35.1457
24.3316 | -35.5375 | -35.1523
24.3377 | -35.5379 | -35.1583
24.3438 | -35.5382 | -35.1642
24.3499 | -35.5386 | -35.1701
24.356 | -35.539 | -35.1761
24.3621 | -35.5394 | -35.1823
24.3682 | -35.5398 | -35.1883
24.3743 | -35.5402 | -35.1944
24.3804 | -35.5406 | -35.1999
24.3865 | -35.541 | -35.2057
24.3926 | -35.5414 | -35.2118
24.3987 | -35.5418 | -35.2179
24.4048 | -35.5422 | -35.2238
24.4109 | -35.5425 | -35.2299
24.417 | -35.5429 | -35.2362
24.4231 | -35.5433 | -35.2423
24.4293 | -35.5437 | -35.2481
24.4354 | -35.5441 | -35.254
24.4415 | -35.5445 | -35.26
24.4476 | -35.545 | -35.266
24.4537 | -35.5453 | -35.2721
24.4598 | -35.5457 | -35.2779
24.4659 | -35.5462 | -35.2841
24.472 | -35.5466 | -35.2902
24.4781 | -35.547 | -35.2964
24.4842 | -35.5474 | -35.3018
24.4903 | -35.5479 | -35.3077
24.4964 | -35.5482 | -35.3137
24.5025 | -35.5486 | -35.3197
24.5086 | -35.549 | -35.3258
24.5147 | -35.5494 | -35.3318
24.5208 | -35.5498 | -35.3375
24.5269 | -35.5501 | -35.3434
24.533 | -35.5506 | -35.3495
24.5391 | -35.551 | -35.3553
24.5452 | -35.5513 | -35.3613
24.5513 | -35.5518 | -35.3669
24.5574 | -35.5522 | -35.3729
24.5635 | -35.5526 | -35.3786
24.5696 | -35.553 | -35.3844
24.5758 | -35.5534 | -35.3904
24.5819 | -35.5538 | -35.3965
24.588 | -35.5542 | -35.4025
24.5941 | -35.5546 | -35.4083
24.6002 | -35.5549 | -35.4143
24.6063 | -35.5553 | -35.4203
24.6124 | -35.5557 | -35.4265
24.6185 | -35.5562 | -35.4324
24.6246 | -35.5565 | -35.4382
24.6307 | -35.5569 | -35.4441
24.6368 | -35.5573 | -35.4503
24.6429 | -35.5577 | -35.4564
24.649 | -35.5581 | -35.4626
24.6551 | -35.5585 | -35.4685
24.6612 | -35.5588 | -35.4744
24.6673 | -35.5592 | -35.4805
24.6734 | -35.5597 | -35.4862
24.6795 | -35.56 | -35.4921
24.6856 | -35.5604 | -35.4984
24.6917 | -35.5608 | -35.5045
24.6978 | -35.5612 | -35.51
24.7039 | -35.5616 | -35.5163
24.71 | -35.562 | -35.5223
24.7162 | -35.5624 | -35.5281
24.7223 | -35.5628 | -35.5341
24.7284 | -35.5632 | -35.5397
24.7345 | -35.5636 | -35.5458
24.7406 | -35.564 | -35.5518
24.7467 | -35.5644 | -35.5579
24.7528 | -35.5648 | -35.5638
24.7589 | -35.5652 | -35.5697
24.765 | -35.5656 | -35.5754
24.7711 | -35.566 | -35.5814
24.7772 | -35.5664 | -35.5872
24.7833 | -35.5668 | -35.5932
24.7894 | -35.5671 | -35.5992
24.7955 | -35.5675 | -35.6052
24.8016 | -35.5678 | -35.611
24.8077 | -35.5682 | -35.6171
24.8138 | -35.5686 | -35.623
24.8199 | -35.569 | -35.6292
24.826 | -35.5694 | -35.635
24.8321 | -35.5697 | -35.6409
24.8382 | -35.5702 | -35.6464
24.8443 | -35.5706 | -35.6525
24.8504 | -35.571 | -35.6585
24.8565 | -35.5714 | -35.6642
24.8627 | -35.5717 | -35.6699
24.8688 | -35.5721 | -35.6757
24.8749 | -35.5725 | -35.6816
24.881 | -35.5728 | -35.6876
24.8871 | -35.5732 | -35.6933
24.8932 | -35.5737 | -35.6991
24.8993 | -35.574 | -35.705
24.9054 | -35.5744 | -35.7106
24.9115 | -35.5748 | -35.7165
24.9176 | -35.5752 | -35.7223
24.9237 | -35.5756 | -35.7281
24.9298 | -35.576 | -35.7341
24.9359 | -35.5764 | -35.74
24.942 | -35.5768 | -35.7455
24.9481 | -35.5772 | -35.7516
24.9542 | -35.5775 | -35.7575
24.9603 | -35.5779 | -35.7634
24.9664 | -35.5783 | -35.7693
24.9725 | -35.5787 | -35.7752
24.9786 | -35.5791 | -35.7806
24.9847 | -35.5794 | -35.7865
24.9908 | -35.5798 | -35.7923
24.9969 | -35.5802 | -35.7983
25.0031 | -35.5806 | -35.8042
25.0092 | -35.581 | -35.8099
25.0153 | -35.5813 | -35.8157
25.0214 | -35.5817 | -35.8216
25.0275 | -35.5821 | -35.8273
25.0336 | -35.5825 | -35.833
25.0397 | -35.5828 | -35.839
25.0458 | -35.5832 | -35.8444
25.0519 | -35.5836 | -35.8504
25.058 | -35.5839 | -35.8562
25.0641 | -35.5843 | -35.8622
25.0702 | -35.5847 | -35.8679
25.0763 | -35.585 | -35.8737
25.0824 | -35.5854 | -35.8792
25.0885 | -35.5858 | -35.8851
25.0946 | -35.5861 | -35.891
25.1007 | -35.5865 | -35.8969
25.1068 | -35.5869 | -35.9026
25.1129 | -35.5872 | -35.9082
25.119 | -35.5876 | -35.9138
25.1251 | -35.588 | -35.9197
25.1312 | -35.5884 | -35.9256
25.1373 | -35.5888 | -35.9314
25.1435 | -35.5892 | -35.9374
25.1496 | -35.5896 | -35.9428
25.1557 | -35.5899 | -35.9485
25.1618 | -35.5903 | -35.9545
25.1679 | -35.5907 | -35.9604
25.174 | -35.5911 | -35.9662
25.1801 | -35.5915 | -35.9721
25.1862 | -35.5918 | -35.9774
25.1923 | -35.5922 | -35.983
25.1984 | -35.5925 | -35.9887
25.2045 | -35.5929 | -35.9943
25.2106 | -35.5933 | -36.0003
25.2167 | -35.5937 | -36.0061
25.2228 | -35.594 | -36.0117
25.2289 | -35.5944 | -36.0175
25.235 | -35.5948 | -36.0234
25.2411 | -35.5952 | -36.0291
25.2472 | -35.5955 | -36.0349
25.2533 | -35.5959 | -36.0409
25.2594 | -35.5962 | -36.0464
25.2655 | -35.5966 | -36.0522
25.2716 | -35.597 | -36.0579
25.2777 | -35.5974 | -36.0635
25.2838 | -35.5978 | -36.0692
25.29 | -35.5981 | -36.0748
25.2961 | -35.5985 | -36.0801
25.3022 | -35.5988 | -36.0857
25.3083 | -35.5992 | -36.0913
25.3144 | -35.5996 | -36.0973
25.3205 | -35.6 | -36.1032
25.3266 | -35.6003 | -36.1085
25.3327 | -35.6007 | -36.1141
25.3388 | -35.6011 | -36.1198
25.3449 | -35.6015 | -36.1257
25.351 | -35.6019 | -36.1314
25.3571 | -35.6023 | -36.1373
25.3632 | -35.6026 | -36.1427
25.3693 | -35.6029 | -36.1485
25.3754 | -35.6033 | -36.1542
25.3815 | -35.6037 | -36.16
25.3876 | -35.6041 | -36.1658
25.3937 | -35.6044 | -36.1716
25.3998 | -35.6048 | -36.1767
25.4059 | -35.6051 | -36.1823
25.412 | -35.6055 | -36.188
25.4181 | -35.6059 | -36.1936
25.4242 | -35.6063 | -36.1995
25.4304 | -35.6066 | -36.2052
25.4365 | -35.607 | -36.2106
25.4426 | -35.6073 | -36.2163
25.4487 | -35.6077 | -36.222
25.4548 | -35.6081 | -36.2276
25.4609 | -35.6085 | -36.2331
25.467 | -35.6088 | -36.2385
25.4731 | -35.6092 | -36.2442
25.4792 | -35.6096 | -36.2497
25.4853 | -35.6099 | -36.2556
25.4914 | -35.6103 | -36.2614
25.4975 | -35.6107 | -36.2671
25.5036 | -35.611 | -36.2727
25.5097 | -35.6113 | -36.2784
25.5158 | -35.6117 | -36.2841
25.5219 | -35.6121 | -36.2899
25.528 | -35.6125 | -36.2957
25.5341 | -35.6129 | -36.3013
25.5402 | -35.6132 | -36.3069
25.5463 | -35.6135 | -36.3124
25.5524 | -35.6139 | -36.3182
25.5585 | -35.6142 | -36.3239
25.5646 | -35.6146 | -36.3296
25.5707 | -35.6149 | -36.3353
25.5769 | -35.6153 | -36.3406
25.583 | -35.6157 | -36.3463
25.5891 | -35.6161 | -36.3521
25.5952 | -35.6164 | -36.3578
25.6013 | -35.6168 | -36.3635
25.6074 | -35.6171 | -36.3687
25.6135 | -35.6175 | -36.3741
25.6196 | -35.6179 | -36.3795
25.6257 | -35.6182 | -36.3852
25.6318 | -35.6186 | -36.391
25.6379 | -35.6189 | -36.3966
25.644 | -35.6193 | -36.4018
25.6501 | -35.6197 | -36.4075
25.6562 | -35.62 | -36.4132
25.6623 | -35.6204 | -36.4189
25.6684 | -35.6207 | -36.4245
25.6745 | -35.6211 | -36.4301
25.6806 | -35.6214 | -36.4356
25.6867 | -35.6218 | -36.4411
25.6928 | -35.6222 | -36.4467
25.6989 | -35.6226 | -36.4525
25.705 | -35.6229 | -36.4583
25.7111 | -35.6233 | -36.4639
25.7173 | -35.6236 | -36.4691
25.7234 | -35.624 | -36.4748
25.7295 | -35.6243 | -36.4803
25.7356 | -35.6247 | -36.4859
25.7417 | -35.625 | -36.4915
25.7478 | -35.6254 | -36.4969
25.7539 | -35.6257 | -36.5025
25.76 | -35.6261 | -36.5081
25.7661 | -35.6264 | -36.5138
25.7722 | -35.6268 | -36.5194
25.7783 | -35.6271 | -36.5249
25.7844 | -35.6274 | -36.5304
25.7905 | -35.6278 | -36.5362
25.7966 | -35.6281 | -36.5419
25.8027 | -35.6285 | -36.5474
25.8088 | -35.6288 | -36.553
25.8149 | -35.6292 | -36.5587
25.821 | -35.6294 | -36.5641
25.8271 | -35.6298 | -36.5698
25.8332 | -35.6301 | -36.5755
25.8393 | -35.6305 | -36.5811
25.8454 | -35.6308 | -36.5867
25.8515 | -35.6312 | -36.5924
25.8576 | -35.6315 | -36.5977
25.8638 | -35.6319 | -36.6033
25.8699 | -35.6322 | -36.6091
25.876 | -35.6326 | -36.6146
25.8821 | -35.6329 | -36.6203
25.8882 | -35.6332 | -36.6255
25.8943 | -35.6335 | -36.631
25.9004 | -35.6339 | -36.6367
25.9065 | -35.6342 | -36.6424
25.9126 | -35.6345 | -36.648
25.9187 | -35.6349 | -36.6534
25.9248 | -35.6352 | -36.6589
25.9309 | -35.6356 | -36.6644
25.937 | -35.6359 | -36.67
25.9431 | -35.6363 | -36.6758
25.9492 | -35.6367 | -36.6813
25.9553 | -35.637 | -36.6871
25.9614 | -35.6373 | -36.6924
25.9675 | -35.6377 | -36.6978
25.9736 | -35.638 | -36.7037
25.9797 | -35.6383 | -36.7095
25.9858 | -35.6387 | -36.7152
25.9919 | -35.639 | -36.7208
25.998 | -35.6394 | -36.7261
26.0042 | -35.6397 | -36.7317
26.0103 | -35.64 | -36.7372
26.0164 | -35.6404 | -36.7425
26.0225 | -35.6407 | -36.7481
26.0286 | -35.641 | -36.7538
26.0347 | -35.6414 | -36.759
26.0408 | -35.6417 | -36.7644
26.0469 | -35.642 | -36.7698
26.053 | -35.6424 | -36.7751
26.0591 | -35.6427 | -36.7806
26.0652 | -35.643 | -36.7861
26.0713 | -35.6434 | -36.7913
26.0774 | -35.6437 | -36.7966
26.0835 | -35.6441 | -36.8024
26.0896 | -35.6444 | -36.8077
26.0957 | -35.6448 | -36.8132
26.1018 | -35.6451 | -36.8185
26.1079 | -35.6454 | -36.8239
26.114 | -35.6457 | -36.8292
26.1201 | -35.6461 | -36.8348
26.1262 | -35.6464 | -36.8402
26.1323 | -35.6468 | -36.8458
26.1384 | -35.6471 | -36.8512
26.1445 | -35.6474 | -36.8566
26.1507 | -35.6478 | -36.8621
26.1568 | -35.6481 | -36.8676
26.1629 | -35.6484 | -36.8732
26.169 | -35.6488 | -36.8786
26.1751 | -35.6491 | -36.8839
26.1812 | -35.6495 | -36.8893
26.1873 | -35.6498 | -36.8947
26.1934 | -35.6501 | -36.9
26.1995 | -35.6505 | -36.9055
26.2056 | -35.6508 | -36.9104
26.2117 | -35.6511 | -36.9161
26.2178 | -35.6514 | -36.9217
26.2239 | -35.6518 | -36.927
26.23 | -35.6521 | -36.9324
26.2361 | -35.6524 | -36.9378
26.2422 | -35.6528 | -36.9429
26.2483 | -35.6531 | -36.9483
26.2544 | -35.6534 | -36.9536
26.2605 | -35.6538 | -36.959
26.2666 | -35.6541 | -36.9644
26.2727 | -35.6544 | -36.9699
26.2788 | -35.6547 | -36.975
26.2849 | -35.6551 | -36.9806
26.2911 | -35.6554 | -36.9859
26.2972 | -35.6558 | -36.9916
26.3033 | -35.6561 | -36.9971
26.3094 | -35.6564 | -37.0024
26.3155 | -35.6567 | -37.0077
26.3216 | -35.6571 | -37.0133
26.3277 | -35.6574 | -37.0185
26.3338 | -35.6577 | -37.0239
26.3399 | -35.658 | -37.0292
26.346 | -35.6583 | -37.0342
26.3521 | -35.6587 | -37.0399
26.3582 | -35.659 | -37.0454
26.3643 | -35.6593 | -37.0506
26.3704 | -35.6596 | -37.0562
26.3765 | -35.66 | -37.0615
26.3826 | -35.6603 | -37.0666
26.3887 | -35.6606 | -37.0719
26.3948 | -35.6609 | -37.0773
26.4009 | -35.6612 | -37.0828
26.407 | -35.6615 | -37.0882
26.4131 | -35.6619 | -37.0937
26.4192 | -35.6622 | -37.0987
26.4253 | -35.6625 | -37.1041
26.4314 | -35.6629 | -37.1097
26.4376 | -35.6632 | -37.1151
26.4437 | -35.6636 | -37.1205
26.4498 | -35.6639 | -37.1261
26.4559 | -35.6642 | -37.1314
26.462 | -35.6645 | -37.1368
26.4681 | -35.6648 | -37.1423
26.4742 | -35.6652 | -37.1477
26.4803 | -35.6655 | -37.1532
26.4864 | -35.6658 | -37.1582
26.4925 | -35.6661 | -37.1637
26.4986 | -35.6665 | -37.1692
26.5047 | -35.6668 | -37.1746
26.5108 | -35.6672 | -37.1799
26.5169 | -35.6675 | -37.1852
26.523 | -35.6678 | -37.1903
26.5291 | -35.6682 | -37.1959
26.5352 | -35.6685 | -37.2014
26.5413 | -35.6688 | -37.2067
26.5474 | -35.6691 | -37.212
26.5535 | -35.6695 | -37.2175
26.5596 | -35.6697 | -37.2227
26.5657 | -35.6701 | -37.2279
26.5718 | -35.6704 | -37.2333
26.578 | -35.6707 | -37.2386
26.5841 | -35.671 | -37.2438
26.5902 | -35.6714 | -37.2494
26.5963 | -35.6717 | -37.2547
26.6024 | -35.672 | -37.26
26.6085 | -35.6723 | -37.2654
26.6146 | -35.6726 | -37.2709
26.6207 | -35.6729 | -37.2762
26.6268 | -35.6733 | -37.2818
26.6329 | -35.6735 | -37.2869
26.639 | -35.6738 | -37.2924
26.6451 | -35.6741 | -37.2978
26.6512 | -35.6745 | -37.303
26.6573 | -35.6748 | -37.3084
26.6634 | -35.6751 | -37.3134
26.6695 | -35.6754 | -37.319
26.6756 | -35.6757 | -37.3246
26.6817 | -35.676 | -37.3301
26.6878 | -35.6763 | -37.3355
26.6939 | -35.6766 | -37.3407
26.7 | -35.6769 | -37.3459
26.7061 | -35.6772 | -37.3511
26.7122 | -35.6775 | -37.3565
26.7183 | -35.6778 | -37.3618
26.7245 | -35.6781 | -37.3671
26.7306 | -35.6785 | -37.3724
26.7367 | -35.6788 | -37.3774
26.7428 | -35.6791 | -37.3827
26.7489 | -35.6794 | -37.3882
26.755 | -35.6797 | -37.3936
26.7611 | -35.68 | -37.399
26.7672 | -35.6803 | -37.4042
26.7733 | -35.6806 | -37.4094
26.7794 | -35.6809 | -37.4149
26.7855 | -35.6812 | -37.4202
26.7916 | -35.6816 | -37.4252
26.7977 | -35.6819 | -37.4307
26.8038 | -35.6822 | -37.4358
26.8099 | -35.6826 | -37.441
26.816 | -35.6829 | -37.4465
26.8221 | -35.6832 | -37.4517
26.8282 | -35.6834 | -37.4569
26.8343 | -35.6837 | -37.4623
26.8404 | -35.684 | -37.4676
26.8465 | -35.6844 | -37.473
26.8526 | -35.6847 | -37.4782
26.8587 | -35.685 | -37.4838
26.8649 | -35.6853 | -37.4889
26.871 | -35.6856 | -37.4942
26.8771 | -35.6859 | -37.4993
26.8832 | -35.6862 | -37.5046
26.8893 | -35.6865 | -37.5099
26.8954 | -35.6868 | -37.5152
26.9015 | -35.6872 | -37.5207
26.9076 | -35.6874 | -37.5261
26.9137 | -35.6877 | -37.5313
26.9198 | -35.688 | -37.5366
26.9259 | -35.6883 | -37.5419
26.932 | -35.6886 | -37.5473
26.9381 | -35.6889 | -37.5527
26.9442 | -35.6892 | -37.558
26.9503 | -35.6895 | -37.5632
26.9564 | -35.6898 | -37.5684
26.9625 | -35.6902 | -37.5739
26.9686 | -35.6905 | -37.5792
26.9747 | -35.6908 | -37.5845
26.9808 | -35.691 | -37.5894
26.9869 | -35.6913 | -37.5948
26.993 | -35.6916 | -37.6001
26.9991 | -35.6919 | -37.6054
27.0052 | -35.6923 | -37.6106
27.0114 | -35.6926 | -37.616
27.0175 | -35.6929 | -37.6211
27.0236 | -35.6932 | -37.6264
27.0297 | -35.6935 | -37.6318
27.0358 | -35.6938 | -37.6371
27.0419 | -35.6941 | -37.6425
27.048 | -35.6945 | -37.6478
27.0541 | -35.6948 | -37.653
27.0602 | -35.6951 | -37.6584
27.0663 | -35.6953 | -37.6636
27.0724 | -35.6956 | -37.6691
27.0785 | -35.696 | -37.6744
27.0846 | -35.6963 | -37.6794
27.0907 | -35.6966 | -37.6849
27.0968 | -35.6969 | -37.6903
27.1029 | -35.6973 | -37.6956
27.109 | -35.6976 | -37.7008
27.1151 | -35.6978 | -37.7062
27.1212 | -35.6981 | -37.7113
27.1273 | -35.6984 | -37.7164
27.1334 | -35.6987 | -37.7217
27.1395 | -35.699 | -37.7271
27.1456 | -35.6993 | -37.7324
27.1518 | -35.6996 | -37.7377
27.1579 | -35.6999 | -37.7426
27.164 | -35.7002 | -37.7481
27.1701 | -35.7005 | -37.7535
27.1762 | -35.7008 | -37.7588
27.1823 | -35.7011 | -37.7642
27.1884 | -35.7014 | -37.7694
27.1945 | -35.7017 | -37.7744
27.2006 | -35.702 | -37.7796
27.2067 | -35.7023 | -37.7848
27.2128 | -35.7026 | -37.7902
27.2189 | -35.7029 | -37.7952
27.225 | -35.7032 | -37.8005
27.2311 | -35.7035 | -37.8056
27.2372 | -35.7038 | -37.8109
27.2433 | -35.7041 | -37.8161
27.2494 | -35.7044 | -37.8215
27.2555 | -35.7047 | -37.8267
27.2616 | -35.705 | -37.8319
27.2677 | -35.7053 | -37.8371
27.2738 | -35.7056 | -37.8422
27.2799 | -35.7059 | -37.8475
27.286 | -35.7062 | -37.8527
27.2921 | -35.7064 | -37.8579
27.2983 | -35.7067 | -37.8629
27.3044 | -35.707 | -37.8683
27.3105 | -35.7073 | -37.8735
27.3166 | -35.7076 | -37.8789
27.3227 | -35.7079 | -37.8843
27.3288 | -35.7082 | -37.8895
27.3349 | -35.7085 | -37.8945
27.341 | -35.7088 | -37.8996
27.3471 | -35.7091 | -37.9047
27.3532 | -35.7094 | -37.9101
27.3593 | -35.7097 | -37.9153
27.3654 | -35.71 | -37.9205
27.3715 | -35.7103 | -37.9255
27.3776 | -35.7106 | -37.9308
27.3837 | -35.7109 | -37.936
27.3898 | -35.7112 | -37.941
27.3959 | -35.7114 | -37.9461
27.402 | -35.7117 | -37.9512
27.4081 | -35.712 | -37.9564
27.4142 | -35.7123 | -37.9617
27.4203 | -35.7127 | -37.9669
27.4264 | -35.713 | -37.972
27.4325 | -35.7133 | -37.9774
27.4387 | -35.7136 | -37.9825
27.4448 | -35.7138 | -37.9876
27.4509 | -35.7142 | -37.9929
27.457 | -35.7145 | -37.9983
27.4631 | -35.7147 | -38.0036
27.4692 | -35.715 | -38.0088
27.4753 | -35.7153 | -38.0141
27.4814 | -35.7156 | -38.0192
27.4875 | -35.7158 | -38.0244
27.4936 | -35.7162 | -38.0295
27.4997 | -35.7165 | -38.0347
27.5058 | -35.7167 | -38.0398
27.5119 | -35.717 | -38.045
27.518 | -35.7173 | -38.0502
27.5241 | -35.7176 | -38.0555
27.5302 | -35.7179 | -38.0609
27.5363 | -35.7181 | -38.0663
27.5424 | -35.7184 | -38.0713
27.5485 | -35.7187 | -38.0765
27.5546 | -35.719 | -38.0817
27.5607 | -35.7193 | -38.087
27.5668 | -35.7196 | -38.0923
27.5729 | -35.7199 | -38.0975
27.5791 | -35.7202 | -38.1027
27.5852 | -35.7205 | -38.1079
27.5913 | -35.7208 | -38.1133
27.5974 | -35.7211 | -38.1185
27.6035 | -35.7214 | -38.1237
27.6096 | -35.7216 | -38.1289
27.6157 | -35.7219 | -38.1339
27.6218 | -35.7222 | -38.1392
27.6279 | -35.7225 | -38.1445
27.634 | -35.7228 | -38.1497
27.6401 | -35.723 | -38.1549
27.6462 | -35.7233 | -38.1603
27.6523 | -35.7235 | -38.1656
27.6584 | -35.7239 | -38.1706
27.6645 | -35.7241 | -38.1758
27.6706 | -35.7244 | -38.1814
27.6767 | -35.7247 | -38.1865
27.6828 | -35.7249 | -38.1915
27.6889 | -35.7252 | -38.1965
27.695 | -35.7254 | -38.2019
27.7011 | -35.7257 | -38.2069
27.7072 | -35.726 | -38.2123
27.7133 | -35.7263 | -38.2176
27.7194 | -35.7266 | -38.2226
27.7256 | -35.7269 | -38.2278
27.7317 | -35.7271 | -38.2331
27.7378 | -35.7274 | -38.2381
27.7439 | -35.7277 | -38.2432
27.75 | -35.728 | -38.2486
27.7561 | -35.7283 | -38.2537
27.7622 | -35.7286 | -38.2589
27.7683 | -35.7289 | -38.2639
27.7744 | -35.7292 | -38.269
27.7805 | -35.7295 | -38.2741
27.7866 | -35.7298 | -38.2792
27.7927 | -35.73 | -38.2843
27.7988 | -35.7303 | -38.2895
27.8049 | -35.7306 | -38.2947
27.811 | -35.7309 | -38.3
27.8171 | -35.7312 | -38.3052
27.8232 | -35.7315 | -38.3102
27.8293 | -35.7318 | -38.3153
27.8354 | -35.7321 | -38.3207
27.8415 | -35.7324 | -38.3261
27.8476 | -35.7326 | -38.3315
27.8537 | -35.7329 | -38.3367
27.8598 | -35.7332 | -38.3418
27.866 | -35.7335 | -38.3468
27.8721 | -35.7337 | -38.3519
27.8782 | -35.734 | -38.3572
27.8843 | -35.7343 | -38.3625
27.8904 | -35.7346 | -38.3678
27.8965 | -35.7348 | -38.3728
27.9026 | -35.7351 | -38.3778
27.9087 | -35.7354 | -38.3827
27.9148 | -35.7357 | -38.3881
27.9209 | -35.736 | -38.3933
27.927 | -35.7363 | -38.3985
27.9331 | -35.7365 | -38.4036
27.9392 | -35.7368 | -38.4088
27.9453 | -35.7371 | -38.414
27.9514 | -35.7373 | -38.4194
27.9575 | -35.7376 | -38.4248
27.9636 | -35.7379 | -38.43
27.9697 | -35.7381 | -38.4352
27.9758 | -35.7384 | -38.4405
27.9819 | -35.7387 | -38.4457
27.988 | -35.739 | -38.451
27.9941 | -35.7392 | -38.4563
28.0002 | -35.7395 | -38.4614
28.0063 | -35.7397 | -38.4667
28.0125 | -35.74 | -38.472
28.0186 | -35.7403 | -38.4771
28.0247 | -35.7406 | -38.482
28.0308 | -35.7408 | -38.4871
28.0369 | -35.7411 | -38.4922
28.043 | -35.7414 | -38.4974
28.0491 | -35.7416 | -38.5026
28.0552 | -35.7419 | -38.5077
28.0613 | -35.7422 | -38.5127
28.0674 | -35.7425 | -38.5179
28.0735 | -35.7427 | -38.523

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4032 | -43.5883
21.4908 | -35.4841 | -35.4726
16.4948 | -34.8778 | -27.9999
12.6602 | -34.5244 | -21.6626
9.7171 | -34.3079 | -16.7112
7.4582 | -34.1584 | -12.7008
5.7244 | -34.0517 | -9.4535
4.3936 | -33.9861 | -6.8377
3.3722 | -33.9742 | -4.7803
2.5883 | -33.9649 | -3.2179
1.9866 | -33.9645 | -1.9738
1.5248 | -33.9646 | -0.97578
1.1703 | -33.9722 | -0.32767
1 | -33.9811 | -0.0061469
0.89822 | -33.9775 | 0.26639
0.68941 | -33.9646 | 0.84506
0.52919 | -33.9665 | 1.2947
0.40612 | -33.9728 | 1.8931
0.31171 | -33.9555 | 2.3425
0.23928 | -33.9527 | 3.0645
0.18358 | -33.9654 | 3.4685
0.14091 | -33.9756 | 4.3443
0.10819 | -33.9933 | 5.4966
0.083017 | -34.0127 | 6.9609
0.063705 | -34.0704 | 8.8204
0.048923 | -34.1267 | 11.304
0.037479 | -34.2331 | 14.4865
0.028801 | -34.4198 | 18.4844
0.022078 | -34.7218 | 23.754
0.017023 | -35.1805 | 30.1173
0.013018 | -35.9648 | 37.7782
0.010014 | -37.0628 | 45.6614

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003673 | -12.7092
0.007347 | -30.0644
0.01102 | -36.226
0.014693 | -41.4442
0.018366 | -44.1605
0.02204 | -45.6361
0.025713 | -48.1551
0.029386 | -50.076
0.033059 | -51.0752
0.036733 | -51.0674
0.040406 | -52.1093
0.044079 | -54.5429
0.047752 | -56.0702
0.051426 | -55.7252
0.055099 | -57.3518
0.058772 | -58.9286
0.062445 | -60.2319
0.066119 | -60.1867
0.069792 | -59.5057
0.073465 | -59.3486
0.077139 | -60.84
0.080812 | -62.0775
0.084485 | -62.6111
0.088158 | -62.8493
0.091832 | -63.4535
0.095505 | -63.8543
0.099178 | -63.5714
0.10285 | -64.4363
0.10652 | -64.4339
0.1102 | -65.2937
0.11387 | -65.0877
0.11754 | -66.0627
0.12122 | -66.4891
0.12489 | -66.4025
0.12856 | -66.9382
0.13224 | -68.2564
0.13591 | -69.1064
0.13958 | -67.5672
0.14326 | -67.2106
0.14693 | -66.8222
0.1506 | -68.0264
0.15428 | -67.5503
0.15795 | -68.1619
0.16162 | -68.0084
0.1653 | -69.2823
0.16897 | -69.8093
0.17264 | -70.7074
0.17632 | -70.2561
0.17999 | -70.7362
0.18366 | -71.0707
0.18734 | -70.9264
0.19101 | -71.6829
0.19468 | -71.0796
0.19836 | -71.2334
0.20203 | -71.1559
0.2057 | -71.4434
0.20938 | -72.4399
0.21305 | -73.4635
0.21672 | -72.722
0.2204 | -72.0723
0.22407 | -72.4512
0.22774 | -72.5886
0.23142 | -72.6224
0.23509 | -72.1071
0.23876 | -72.1969
0.24244 | -73.2224
0.24611 | -73.6179
0.24978 | -73.9648
0.25345 | -73.36
0.25713 | -73.5181
0.2608 | -73.2629
0.26448 | -73.7873
0.26815 | -74.53
0.27182 | -75.109
0.27549 | -74.1022
0.27917 | -73.4062
0.28284 | -73.5153
0.28652 | -74.4899
0.29019 | -74.6353
0.29386 | -75.7314
0.29753 | -75.844
0.30121 | -75.9164
0.30488 | -75.7189
0.30855 | -75.1888
0.31223 | -75.9403
0.3159 | -76.1039
0.31957 | -76.1312
0.32325 | -75.658
0.32692 | -75.8309
0.33059 | -76.6783
0.33427 | -77.2431
0.33794 | -77.0327
0.34161 | -76.104
0.34529 | -76.2041
0.34896 | -77.0989
0.35263 | -77.8499
0.35631 | -78.3353
0.35998 | -77.5003
0.36365 | -77.665
0.36733 | -76.8207
0.371 | -76.4832
0.37467 | -77.4956
0.37835 | -77.9601
0.38202 | -77.0707
0.38569 | -76.9199
0.38937 | -77.3496
0.39304 | -78.1925
0.39671 | -78.5279
0.40039 | -79.1368
0.40406 | -78.1187
0.40773 | -78.4567
0.41141 | -77.7612
0.41508 | -78.5403
0.41875 | -78.7514
0.42242 | -79.3992
0.4261 | -78.9541
0.42977 | -78.4959
0.43345 | -78.7941
0.43712 | -78.6052
0.44079 | -79.0643
0.44446 | -80.3862
0.44814 | -80.4436
0.45181 | -79.6892
0.45548 | -80.281
0.45916 | -78.8273
0.46283 | -78.3153
0.4665 | -79.7689
0.47018 | -79.6874
0.47385 | -79.3505
0.47752 | -79.5468
0.4812 | -78.9093
0.48487 | -79.0699
0.48854 | -80.3039
0.49222 | -81.2704
0.49589 | -80.6293
0.49956 | -79.7445
0.50324 | -80.1247
0.50691 | -81.0212
0.51058 | -80.9852
0.51426 | -80.1769
0.51793 | -80.9476
0.5216 | -81.4107
0.52528 | -80.5925
0.52895 | -80.5479
0.53262 | -80.4995
0.5363 | -80.6193
0.53997 | -81.3681
0.54364 | -81.7861
0.54732 | -81.2818
0.55099 | -81.0524
0.55466 | -81.2555
0.55834 | -81.4396
0.56201 | -81.3988
0.56568 | -81.3993
0.56936 | -81.6727
0.57303 | -82.0696
0.5767 | -81.7278
0.58038 | -81.4561
0.58405 | -81.7811
0.58772 | -81.8928
0.5914 | -81.8811
0.59507 | -82.1797
0.59874 | -82.6239
0.60242 | -82.6666
0.60609 | -82.3903
0.60976 | -82.2768
0.61343 | -82.0063
0.61711 | -81.7514
0.62078 | -81.7535
0.62445 | -81.8658
0.62813 | -82.2653
0.6318 | -82.7832
0.63548 | -82.7316
0.63915 | -82.7345
0.64282 | -82.7115
0.64649 | -82.869
0.65017 | -83.1669
0.65384 | -83.4158
0.65751 | -83.3834
0.66119 | -83.1817
0.66486 | -83.0951
0.66853 | -83.2445
0.67221 | -83.5879
0.67588 | -83.8391
0.67955 | -83.7096
0.68323 | -83.6865
0.6869 | -83.5533
0.69057 | -83.2633
0.69425 | -83.2412
0.69792 | -83.2353
0.70159 | -83.0495
0.70527 | -82.9072
0.70894 | -83.1528
0.71261 | -83.6584
0.71629 | -83.5889
0.71996 | -83.2494
0.72363 | -83.4187
0.72731 | -83.8486
0.73098 | -83.8995
0.73465 | -83.9038
0.73833 | -83.947
0.742 | -83.9523
0.74567 | -84.1463
0.74935 | -84.243
0.75302 | -84.185
0.75669 | -84.2577
0.76036 | -84.235
0.76404 | -84.1291
0.76771 | -83.7949
0.77138 | -83.1727
0.77506 | -83.2385
0.77873 | -84.2707
0.78241 | -85.283
0.78608 | -85.3355
0.78975 | -84.8772
0.79343 | -84.5614
0.7971 | -84.3914
0.80077 | -84.1509
0.80444 | -84.0258
0.80812 | -84.1883
0.81179 | -84.461
0.81546 | -84.7068
0.81914 | -85.2606
0.82281 | -85.6406
0.82648 | -85.4099
0.83016 | -85.0467
0.83383 | -84.8035
0.8375 | -84.6497
0.84118 | -84.6958
0.84485 | -84.9193
0.84852 | -85.2961
0.8522 | -85.4901
0.85587 | -85.112
0.85954 | -84.9868
0.86322 | -85.4113
0.86689 | -85.5568
0.87056 | -85.4218
0.87424 | -85.067
0.87791 | -84.7578
0.88158 | -85.0268
0.88526 | -85.6016
0.88893 | -86.0358
0.8926 | -85.9288
0.89628 | -85.511
0.89995 | -85.2634
0.90362 | -85.3515
0.9073 | -85.4694
0.91097 | -85.5573
0.91464 | -85.4102
0.91832 | -85.2493
0.92199 | -85.3998
0.92566 | -85.8166
0.92934 | -86.2837
0.93301 | -86.1283
0.93668 | -85.6473
0.94036 | -85.5129
0.94403 | -85.9104
0.9477 | -86.4597
0.95137 | -86.6803
0.95505 | -87.1131
0.95872 | -87.559
0.9624 | -87.2024
0.96607 | -86.5702
0.96974 | -86.0495
0.97342 | -85.9033
0.97709 | -86.0734
0.98076 | -86.0317
0.98443 | -85.8863
0.98811 | -86.074
0.99178 | -86.4736
0.99545 | -86.6753
0.99913 | -86.6351
1.0028 | -86.7046
1.0065 | -86.8593
1.0101 | -87.1329
1.0138 | -87.4577
1.0175 | -87.2498
1.0212 | -86.839
1.0248 | -86.8919
1.0285 | -87.4951
1.0322 | -87.7909
1.0359 | -87.6833
1.0395 | -87.4349
1.0432 | -87.3377
1.0469 | -87.5022
1.0506 | -87.4687
1.0542 | -87.3281
1.0579 | -87.4158
1.0616 | -87.388
1.0652 | -87.3228
1.0689 | -87.4152
1.0726 | -87.7412
1.0763 | -88.0919
1.0799 | -88.2451
1.0836 | -88.2456
1.0873 | -88.354
1.091 | -88.6765
1.0946 | -88.8973
1.0983 | -88.7508
1.102 | -88.3763
1.1057 | -88.004
1.1093 | -87.8361
1.113 | -87.9362
1.1167 | -88.1587
1.1203 | -88.3711
1.124 | -88.498
1.1277 | -88.5318
1.1314 | -88.6511
1.135 | -88.9394
1.1387 | -89.1543
1.1424 | -89.001
1.1461 | -88.5182
1.1497 | -88.0577
1.1534 | -87.9163
1.1571 | -88.0567
1.1608 | -88.2308
1.1644 | -88.4291
1.1681 | -88.7092
1.1718 | -88.8098
1.1754 | -88.6004
1.1791 | -88.313
1.1828 | -88.2299
1.1865 | -88.4924
1.1901 | -89.0165
1.1938 | -89.5156
1.1975 | -89.6567
1.2012 | -89.401
1.2048 | -89.0412
1.2085 | -88.8052
1.2122 | -88.7736
1.2159 | -88.8572
1.2195 | -88.7986
1.2232 | -88.4903
1.2269 | -88.2083
1.2305 | -88.2268
1.2342 | -88.4566
1.2379 | -88.64
1.2416 | -88.7495
1.2452 | -88.9009
1.2489 | -89.028
1.2526 | -88.9423
1.2563 | -88.754
1.2599 | -88.7056
1.2636 | -88.7069
1.2673 | -88.5624
1.2709 | -88.345
1.2746 | -88.2857
1.2783 | -88.4573
1.282 | -88.6766
1.2856 | -88.8013
1.2893 | -88.8787
1.293 | -89.0305
1.2967 | -89.3456
1.3003 | -89.7426
1.304 | -89.9957
1.3077 | -89.9468
1.3114 | -89.6597
1.315 | -89.2826
1.3187 | -89.0901
1.3224 | -89.3393
1.326 | -89.8141
1.3297 | -90.0327
1.3334 | -89.9202
1.3371 | -89.6847
1.3407 | -89.3921
1.3444 | -89.1344
1.3481 | -89.1993
1.3518 | -89.6677
1.3554 | -90.1192
1.3591 | -90.1507
1.3628 | -89.979
1.3665 | -89.8885
1.3701 | -89.8312
1.3738 | -89.7096
1.3775 | -89.5922
1.3811 | -89.6109
1.3848 | -89.7558
1.3885 | -89.888
1.3922 | -89.827
1.3958 | -89.4993
1.3995 | -89.1547
1.4032 | -89.0513
1.4069 | -89.1489
1.4105 | -89.3083
1.4142 | -89.4609
1.4179 | -89.5699
1.4216 | -89.6547
1.4252 | -89.7947
1.4289 | -89.9958
1.4326 | -90.1595
1.4362 | -90.2547
1.4399 | -90.3218
1.4436 | -90.2873
1.4473 | -90.1802
1.4509 | -90.1519
1.4546 | -90.173
1.4583 | -90.1722
1.462 | -90.2202
1.4656 | -90.4099
1.4693 | -90.7508
1.473 | -91.1187
1.4767 | -91.306
1.4803 | -91.1894
1.484 | -90.8317
1.4877 | -90.4324
1.4913 | -90.2576
1.495 | -90.3507
1.4987 | -90.4861
1.5024 | -90.5006
1.506 | -90.4124
1.5097 | -90.2998
1.5134 | -90.3196
1.5171 | -90.5161
1.5207 | -90.6952
1.5244 | -90.7332
1.5281 | -90.8362
1.5318 | -91.027
1.5354 | -91.0184
1.5391 | -90.7962
1.5428 | -90.6175
1.5464 | -90.5021
1.5501 | -90.3311
1.5538 | -90.194
1.5575 | -90.2057
1.5611 | -90.3651
1.5648 | -90.6505
1.5685 | -90.8845
1.5722 | -90.8399
1.5758 | -90.7045
1.5795 | -90.7624
1.5832 | -90.9202
1.5869 | -91.0775
1.5905 | -91.2865
1.5942 | -91.5068
1.5979 | -91.5967
1.6015 | -91.4988
1.6052 | -91.3226
1.6089 | -91.197
1.6126 | -91.1737
1.6162 | -91.2553
1.6199 | -91.3786
1.6236 | -91.4217
1.6273 | -91.3317
1.6309 | -91.1923
1.6346 | -91.1236
1.6383 | -91.1532
1.6419 | -91.2209
1.6456 | -91.3016
1.6493 | -91.4358
1.653 | -91.6658
1.6566 | -91.9341
1.6603 | -92.1106
1.664 | -92.0724
1.6677 | -91.826
1.6713 | -91.5155
1.675 | -91.2946
1.6787 | -91.2026
1.6824 | -91.1909
1.686 | -91.2123
1.6897 | -91.2517
1.6934 | -91.2999
1.697 | -91.3766
1.7007 | -91.5042
1.7044 | -91.6776
1.7081 | -91.8275
1.7117 | -91.869
1.7154 | -91.8085
1.7191 | -91.7467
1.7228 | -91.7465
1.7264 | -91.7608
1.7301 | -91.7164
1.7338 | -91.6417
1.7375 | -91.621
1.7411 | -91.6823
1.7448 | -91.796
1.7485 | -91.9431
1.7521 | -92.0335
1.7558 | -91.9471
1.7595 | -91.6552
1.7632 | -91.3123
1.7668 | -91.1593
1.7705 | -91.3437
1.7742 | -91.8286
1.7779 | -92.3924
1.7815 | -92.7383
1.7852 | -92.7285
1.7889 | -92.4301
1.7926 | -92.0527
1.7962 | -91.7889
1.7999 | -91.6973
1.8036 | -91.7096
1.8072 | -91.759
1.8109 | -91.8449
1.8146 | -91.9948
1.8183 | -92.1919
1.8219 | -92.3926
1.8256 | -92.5703
1.8293 | -92.6762
1.833 | -92.6376
1.8366 | -92.4817
1.8403 | -92.3203
1.844 | -92.2139
1.8477 | -92.136
1.8513 | -92.041
1.855 | -91.9246
1.8587 | -91.8314
1.8623 | -91.8106
1.866 | -91.8756
1.8697 | -91.9597
1.8734 | -92.0205
1.877 | -92.1341
1.8807 | -92.3777
1.8844 | -92.6887
1.8881 | -92.8489
1.8917 | -92.718
1.8954 | -92.3959
1.8991 | -92.1011
1.9027 | -91.9674
1.9064 | -91.9809
1.9101 | -92.0478
1.9138 | -92.1473
1.9174 | -92.3353
1.9211 | -92.6214
1.9248 | -92.9182
1.9285 | -93.1336
1.9321 | -93.2447
1.9358 | -93.242
1.9395 | -93.091
1.9432 | -92.8281
1.9468 | -92.5763
1.9505 | -92.4959
1.9542 | -92.6703
1.9578 | -93.0267
1.9615 | -93.3473
1.9652 | -93.3599
1.9689 | -93.0245
1.9725 | -92.5717
1.9762 | -92.2576
1.9799 | -92.17
1.9836 | -92.2426
1.9872 | -92.3589
1.9909 | -92.4483
1.9946 | -92.4996
1.9983 | -92.5078
2.0019 | -92.4753
2.0056 | -92.443
2.0093 | -92.4561
2.0129 | -92.5234
2.0166 | -92.6119
2.0203 | -92.6319
2.024 | -92.5376
2.0276 | -92.4496
2.0313 | -92.518
2.035 | -92.7755
2.0387 | -93.144
2.0423 | -93.5015
2.046 | -93.7389
2.0497 | -93.7856
2.0534 | -93.683
2.057 | -93.5163
2.0607 | -93.3545
2.0644 | -93.2918
2.068 | -93.3854
2.0717 | -93.6368
2.0754 | -93.962
2.0791 | -94.2175
2.0827 | -94.3064
2.0864 | -94.2467
2.0901 | -94.0795
2.0938 | -93.8574
2.0974 | -93.6421
2.1011 | -93.4531
2.1048 | -93.3026
2.1085 | -93.2231
2.1121 | -93.2333
2.1158 | -93.3132
2.1195 | -93.3804
2.1231 | -93.3825
2.1268 | -93.3232
2.1305 | -93.2243
2.1342 | -93.1508
2.1378 | -93.1556
2.1415 | -93.2162
2.1452 | -93.2829
2.1489 | -93.3213
2.1525 | -93.3234
2.1562 | -93.2969
2.1599 | -93.2667
2.1636 | -93.22
2.1672 | -93.1388
2.1709 | -93.0616
2.1746 | -92.9987
2.1782 | -92.9753
2.1819 | -92.9927
2.1856 | -93.0258
2.1893 | -93.0739
2.1929 | -93.1245
2.1966 | -93.1774
2.2003 | -93.2728
2.204 | -93.455
2.2076 | -93.7064
2.2113 | -93.9504
2.215 | -94.1216
2.2187 | -94.117
2.2223 | -93.9266
2.226 | -93.6589
2.2297 | -93.411
2.2333 | -93.2596
2.237 | -93.2161
2.2407 | -93.2482
2.2444 | -93.3146
2.248 | -93.3612
2.2517 | -93.3361
2.2554 | -93.2491
2.2591 | -93.1667
2.2627 | -93.1205
2.2664 | -93.109
2.2701 | -93.1345
2.2738 | -93.1952
2.2774 | -93.2808
2.2811 | -93.3863
2.2848 | -93.503
2.2884 | -93.6263
2.2921 | -93.7707
2.2958 | -93.9349
2.2995 | -94.1215
2.3031 | -94.3103
2.3068 | -94.4498
2.3105 | -94.4963
2.3142 | -94.4542
2.3178 | -94.378
2.3215 | -94.2839
2.3252 | -94.153
2.3288 | -93.9911
2.3325 | -93.8233
2.3362 | -93.7192
2.3399 | -93.7007
2.3435 | -93.7524
2.3472 | -93.8363
2.3509 | -93.9261
2.3546 | -94.0083
2.3582 | -94.0675
2.3619 | -94.0907
2.3656 | -94.0799
2.3693 | -94.0434
2.3729 | -94.0042
2.3766 | -93.968
2.3803 | -93.9165
2.3839 | -93.8401
2.3876 | -93.7621
2.3913 | -93.7356
2.395 | -93.7867
2.3986 | -93.9094
2.4023 | -94.0838
2.406 | -94.2713
2.4097 | -94.4458
2.4133 | -94.5986
2.417 | -94.7036
2.4207 | -94.7282
2.4244 | -94.6422
2.428 | -94.4696
2.4317 | -94.2881
2.4354 | -94.1892
2.439 | -94.2038
2.4427 | -94.2864
2.4464 | -94.3775
2.4501 | -94.4137
2.4537 | -94.3652
2.4574 | -94.2373
2.4611 | -94.0557
2.4648 | -93.8704
2.4684 | -93.7609
2.4721 | -93.775
2.4758 | -93.9166
2.4795 | -94.1487
2.4831 | -94.4043
2.4868 | -94.621
2.4905 | -94.7821
2.4941 | -94.8611
2.4978 | -94.8769
2.5015 | -94.851
2.5052 | -94.8163
2.5088 | -94.7872
2.5125 | -94.7593
2.5162 | -94.7181
2.5199 | -94.6568
2.5235 | -94.6146
2.5272 | -94.6036
2.5309 | -94.631
2.5346 | -94.7082
2.5382 | -94.8074
2.5419 | -94.9166
2.5456 | -94.9652
2.5492 | -94.9116
2.5529 | -94.7861
2.5566 | -94.673
2.5603 | -94.6418
2.5639 | -94.6849
2.5676 | -94.7612
2.5713 | -94.8241
2.575 | -94.8558
2.5786 | -94.8681
2.5823 | -94.8684
2.586 | -94.8842
2.5897 | -94.93
2.5933 | -94.9463
2.597 | -94.887
2.6007 | -94.7807
2.6043 | -94.704
2.608 | -94.7473
2.6117 | -94.9244
2.6154 | -95.1681
2.619 | -95.3704
2.6227 | -95.4209
2.6264 | -95.3046
2.6301 | -95.109
2.6337 | -94.9546
2.6374 | -94.9002
2.6411 | -94.9621
2.6448 | -95.1181
2.6484 | -95.2669
2.6521 | -95.3514
2.6558 | -95.3497
2.6594 | -95.2348
2.6631 | -95.0275
2.6668 | -94.7934
2.6705 | -94.5949
2.6741 | -94.4732
2.6778 | -94.4601
2.6815 | -94.531
2.6852 | -94.6466
2.6888 | -94.7745
2.6925 | -94.8896
2.6962 | -94.9872
2.6998 | -95.0843
2.7035 | -95.1723
2.7072 | -95.2584
2.7109 | -95.3345
2.7145 | -95.4164
2.7182 | -95.493
2.7219 | -95.5737
2.7256 | -95.6526
2.7292 | -95.7176
2.7329 | -95.7507
2.7366 | -95.7367
2.7403 | -95.6905
2.7439 | -95.6305
2.7476 | -95.5668
2.7513 | -95.5172
2.7549 | -95.4672
2.7586 | -95.4202
2.7623 | -95.3887
2.766 | -95.3883
2.7696 | -95.4043
2.7733 | -95.4216
2.777 | -95.4394
2.7807 | -95.4392
2.7843 | -95.4185
2.788 | -95.3824
2.7917 | -95.3358
2.7954 | -95.3073
2.799 | -95.3092
2.8027 | -95.3367
2.8064 | -95.3684
2.81 | -95.3997
2.8137 | -95.4391
2.8174 | -95.5003
2.8211 | -95.5893
2.8247 | -95.704
2.8284 | -95.8445
2.8321 | -95.993
2.8358 | -96.1318
2.8394 | -96.2488
2.8431 | -96.3171
2.8468 | -96.3202
2.8505 | -96.2695
2.8541 | -96.176
2.8578 | -96.0422
2.8615 | -95.9021
2.8651 | -95.7757
2.8688 | -95.6758
2.8725 | -95.6101
2.8762 | -95.5823
2.8798 | -95.5872
2.8835 | -95.6054
2.8872 | -95.6332
2.8909 | -95.6492
2.8945 | -95.6505
2.8982 | -95.6403
2.9019 | -95.6295
2.9056 | -95.6396
2.9092 | -95.6674
2.9129 | -95.709
2.9166 | -95.7559
2.9202 | -95.8184
2.9239 | -95.9099
2.9276 | -96.0231
2.9313 | -96.1081
2.9349 | -96.0983
2.9386 | -95.9825
2.9423 | -95.795
2.946 | -95.604
2.9496 | -95.4559
2.9533 | -95.3756
2.957 | -95.3712
2.9606 | -95.4344
2.9643 | -95.5437
2.968 | -95.6779
2.9717 | -95.8247
2.9753 | -95.9623
2.979 | -96.0849
2.9827 | -96.1964
2.9864 | -96.2853
2.99 | -96.3414
2.9937 | -96.3672
2.9974 | -96.3453
3.0011 | -96.2594
3.0047 | -96.1486
3.0084 | -96.0431
3.0121 | -95.9786
3.0157 | -95.954
3.0194 | -95.962
3.0231 | -96.0171
3.0268 | -96.1241
3.0304 | -96.2864
3.0341 | -96.4911
3.0378 | -96.6929
3.0415 | -96.8439
3.0451 | -96.9386
3.0488 | -96.9409
3.0525 | -96.8303
3.0562 | -96.6482
3.0598 | -96.4029
3.0635 | -96.1277
3.0672 | -95.8584
3.0708 | -95.6185
3.0745 | -95.4365
3.0782 | -95.3335
3.0819 | -95.3045
3.0855 | -95.3474
3.0892 | -95.4514
3.0929 | -95.5983
3.0966 | -95.786
3.1002 | -95.9724
3.1039 | -96.1236
3.1076 | -96.2285
3.1113 | -96.2717
3.1149 | -96.2723
3.1186 | -96.2674
3.1223 | -96.2937
3.1259 | -96.3472
3.1296 | -96.4276
3.1333 | -96.5008
3.137 | -96.5424
3.1406 | -96.5582
3.1443 | -96.5489
3.148 | -96.5389
3.1517 | -96.5339
3.1553 | -96.5485
3.159 | -96.5895
3.1627 | -96.6554
3.1664 | -96.7298
3.17 | -96.784
3.1737 | -96.8027
3.1774 | -96.7481
3.181 | -96.7209
3.1847 | -96.6916
3.1884 | -96.6915
3.1921 | -96.7213
3.1957 | -96.7644
3.1994 | -96.8006
3.2031 | -96.8057
3.2068 | -96.7653
3.2104 | -96.6811
3.2141 | -96.5708
3.2178 | -96.4627
3.2215 | -96.3783
3.2251 | -96.3242
3.2288 | -96.3101
3.2325 | -96.331
3.2361 | -96.37
3.2398 | -96.4087
3.2435 | -96.4284
3.2472 | -96.4302
3.2508 | -96.4381
3.2545 | -96.4601
3.2582 | -96.4936
3.2619 | -96.5431
3.2655 | -96.6036
3.2692 | -96.667
3.2729 | -96.7334
3.2766 | -96.8
3.2802 | -96.8649
3.2839 | -96.8959
3.2876 | -96.887
3.2912 | -96.8422
3.2949 | -96.7632
3.2986 | -96.6829
3.3023 | -96.6201
3.3059 | -96.5774
3.3096 | -96.5647
3.3133 | -96.591
3.317 | -96.6479
3.3206 | -96.7133
3.3243 | -96.791
3.328 | -96.867
3.3316 | -96.9316
3.3353 | -96.9593
3.339 | -96.9679
3.3427 | -96.9561
3.3463 | -96.9214
3.35 | -96.8975
3.3537 | -96.8938
3.3574 | -96.9321
3.361 | -97.0166
3.3647 | -97.1501
3.3684 | -97.3241
3.3721 | -97.5153
3.3757 | -97.6938
3.3794 | -97.8032
3.3831 | -97.8284
3.3867 | -97.7489
3.3904 | -97.5891
3.3941 | -97.3564
3.3978 | -97.1001
3.4014 | -96.8812
3.4051 | -96.7303
3.4088 | -96.6624
3.4125 | -96.6708
3.4161 | -96.7461
3.4198 | -96.8695
3.4235 | -97.0273
3.4272 | -97.2079
3.4308 | -97.4014
3.4345 | -97.5877
3.4382 | -97.7351
3.4418 | -97.8426
3.4455 | -97.8847
3.4492 | -97.8769
3.4529 | -97.838
3.4565 | -97.7712
3.4602 | -97.6878
3.4639 | -97.6139
3.4676 | -97.5478
3.4712 | -97.4766
3.4749 | -97.4055
3.4786 | -97.3401
3.4823 | -97.3021
3.4859 | -97.291
3.4896 | -97.3099
3.4933 | -97.3506
3.4969 | -97.3928
3.5006 | -97.4359
3.5043 | -97.4561
3.508 | -97.4463
3.5116 | -97.4245
3.5153 | -97.392
3.519 | -97.3387
3.5227 | -97.2762
3.5263 | -97.2197
3.53 | -97.1691
3.5337 | -97.1547
3.5374 | -97.193
3.541 | -97.2974
3.5447 | -97.4483
3.5484 | -97.6252
3.552 | -97.7956
3.5557 | -97.9249
3.5594 | -97.9953
3.5631 | -97.9971
3.5667 | -97.9417
3.5704 | -97.8775
3.5741 | -97.8383
3.5778 | -97.8255
3.5814 | -97.8347
3.5851 | -97.8723
3.5888 | -97.9121
3.5925 | -97.9395
3.5961 | -97.9492
3.5998 | -97.9531
3.6035 | -97.9609
3.6071 | -97.9747
3.6108 | -98.007
3.6145 | -98.051
3.6182 | -98.0904
3.6218 | -98.1084
3.6255 | -98.0763
3.6292 | -98.0008
3.6329 | -97.9083
3.6365 | -97.7995
3.6402 | -97.6852
3.6439 | -97.5906
3.6475 | -97.5263
3.6512 | -97.4997
3.6549 | -97.502
3.6586 | -97.5308
3.6622 | -97.5897
3.6659 | -97.6549
3.6696 | -97.7233
3.6733 | -97.7918
3.6769 | -97.8426
3.6806 | -97.8829
3.6843 | -97.9053
3.688 | -97.9054
3.6916 | -97.8993
3.6953 | -97.8852
3.699 | -97.8576
3.7026 | -97.8304
3.7063 | -97.8217
3.71 | -97.8035
3.7137 | -97.7927
3.7173 | -97.7866
3.721 | -97.7766
3.7247 | -97.7729
3.7284 | -97.7691
3.732 | -97.7782
3.7357 | -97.8056
3.7394 | -97.8548
3.7431 | -97.9257
3.7467 | -98.0232
3.7504 | -98.1367
3.7541 | -98.2637
3.7577 | -98.3982
3.7614 | -98.5201
3.7651 | -98.6167
3.7688 | -98.6795
3.7724 | -98.7072
3.7761 | -98.6954
3.7798 | -98.6472
3.7835 | -98.5878
3.7871 | -98.534
3.7908 | -98.488
3.7945 | -98.4388
3.7982 | -98.4044
3.8018 | -98.3686
3.8055 | -98.3428
3.8092 | -98.3177
3.8128 | -98.2844
3.8165 | -98.2442
3.8202 | -98.1941
3.8239 | -98.1412
3.8275 | -98.0917
3.8312 | -98.0623
3.8349 | -98.0544
3.8386 | -98.075
3.8422 | -98.1265
3.8459 | -98.1943
3.8496 | -98.2671
3.8533 | -98.347
3.8569 | -98.4292
3.8606 | -98.4999
3.8643 | -98.5627
3.8679 | -98.6044
3.8716 | -98.6201
3.8753 | -98.6004
3.879 | -98.5411
3.8826 | -98.4454
3.8863 | -98.3202
3.89 | -98.1809
3.8937 | -98.044
3.8973 | -97.9233
3.901 | -97.8326
3.9047 | -97.7789
3.9084 | -97.7564
3.912 | -97.7595
3.9157 | -97.795
3.9194 | -97.8558
3.923 | -97.9382
3.9267 | -98.0447
3.9304 | -98.1577
3.9341 | -98.2706
3.9377 | -98.3788
3.9414 | -98.4627
3.9451 | -98.5224
3.9488 | -98.5624
3.9524 | -98.5942
3.9561 | -98.6133
3.9598 | -98.6333
3.9635 | -98.6614
3.9671 | -98.7027
3.9708 | -98.75
3.9745 | -98.8001
3.9781 | -98.8451
3.9818 | -98.8903
3.9855 | -98.9212
3.9892 | -98.9213
3.9928 | -98.9059
3.9965 | -98.866
4.0002 | -98.7959
4.0039 | -98.7213
4.0075 | -98.6487
4.0112 | -98.5914
4.0149 | -98.5596
4.0186 | -98.5661
4.0222 | -98.5987
4.0259 | -98.6494
4.0296 | -98.7042
4.0332 | -98.7776
4.0369 | -98.8516
4.0406 | -98.9074
4.0443 | -98.9453
4.0479 | -98.9572
4.0516 | -98.9293
4.0553 | -98.8705
4.059 | -98.7947
4.0626 | -98.7201
4.0663 | -98.6527
4.07 | -98.605
4.0736 | -98.5846
4.0773 | -98.5841
4.081 | -98.6046
4.0847 | -98.6435
4.0883 | -98.6941
4.092 | -98.7556
4.0957 | -98.7997
4.0994 | -98.837
4.103 | -98.869
4.1067 | -98.8869
4.1104 | -98.9006
4.1141 | -98.9058
4.1177 | -98.8863
4.1214 | -98.8535
4.1251 | -98.8078
4.1287 | -98.7537
4.1324 | -98.6949
4.1361 | -98.6414
4.1398 | -98.6041
4.1434 | -98.5793
4.1471 | -98.5759
4.1508 | -98.5994
4.1545 | -98.6442
4.1581 | -98.7071
4.1618 | -98.7695
4.1655 | -98.8166
4.1692 | -98.85
4.1728 | -98.855
4.1765 | -98.8411
4.1802 | -98.8286
4.1838 | -98.8202
4.1875 | -98.8338
4.1912 | -98.8723
4.1949 | -98.9318
4.1985 | -99.003
4.2022 | -99.0813
4.2059 | -99.1553
4.2096 | -99.1985
4.2132 | -99.2098
4.2169 | -99.1968
4.2206 | -99.1601
4.2243 | -99.1032
4.2279 | -99.0457
4.2316 | -98.9993
4.2353 | -98.9796
4.2389 | -99.0116
4.2426 | -99.0249
4.2463 | -99.0437
4.25 | -99.0707
4.2536 | -99.0911
4.2573 | -99.1091
4.261 | -99.1294
4.2647 | -99.1383
4.2683 | -99.1366
4.272 | -99.1237
4.2757 | -99.1065
4.2794 | -99.0873
4.283 | -99.0723
4.2867 | -99.0755
4.2904 | -99.1087
4.294 | -99.1655
4.2977 | -99.2396
4.3014 | -99.3166
4.3051 | -99.392
4.3087 | -99.4643
4.3124 | -99.5267
4.3161 | -99.5689
4.3198 | -99.5989
4.3234 | -99.6225
4.3271 | -99.6344
4.3308 | -99.6362
4.3345 | -99.6403
4.3381 | -99.6528
4.3418 | -99.6697
4.3455 | -99.6814
4.3491 | -99.6889
4.3528 | -99.6901
4.3565 | -99.6702
4.3602 | -99.6264
4.3638 | -99.5713
4.3675 | -99.5081
4.3712 | -99.4398
4.3749 | -99.3716
4.3785 | -99.3018
4.3822 | -99.2359
4.3859 | -99.1861
4.3895 | -99.1469
4.3932 | -99.1261
4.3969 | -99.1318
4.4006 | -99.1655
4.4042 | -99.2194
4.4079 | -99.2839
4.4116 | -99.3616
4.4153 | -99.4544
4.4189 | -99.5397
4.4226 | -99.6087
4.4263 | -99.6608
4.43 | -99.6814
4.4336 | -99.6618
4.4373 | -99.6245
4.441 | -99.5844
4.4446 | -99.5463
4.4483 | -99.5268
4.452 | -99.5234
4.4557 | -99.5405
4.4593 | -99.5797
4.463 | -99.6384
4.4667 | -99.7127
4.4704 | -99.7931
4.474 | -99.8776
4.4777 | -99.964
4.4814 | -100.0432
4.4851 | -100.1059
4.4887 | -100.1581
4.4924 | -100.1977
4.4961 | -100.2197
4.4997 | -100.2364
4.5034 | -100.2466
4.5071 | -100.2384
4.5108 | -100.224
4.5144 | -100.2031
4.5181 | -100.1679
4.5218 | -100.1206
4.5255 | -100.0321
4.5291 | -99.9258
4.5328 | -99.8056
4.5365 | -99.6799
4.5402 | -99.5654
4.5438 | -99.4684
4.5475 | -99.3891
4.5512 | -99.3326
4.5548 | -99.3016
4.5585 | -99.2988
4.5622 | -99.3215
4.5659 | -99.3604
4.5695 | -99.4047
4.5732 | -99.4617
4.5769 | -99.5198
4.5806 | -99.5722
4.5842 | -99.6182
4.5879 | -99.6557
4.5916 | -99.6796
4.5953 | -99.7021
4.5989 | -99.7164
4.6026 | -99.7316
4.6063 | -99.751
4.6099 | -99.7785
4.6136 | -99.8077
4.6173 | -99.8396
4.621 | -99.8605
4.6246 | -99.8671
4.6283 | -99.8664
4.632 | -99.8669
4.6357 | -99.8704
4.6393 | -99.8854
4.643 | -99.9089
4.6467 | -99.9415
4.6504 | -99.9748
4.654 | -100.0184
4.6577 | -100.0705
4.6614 | -100.1296
4.665 | -100.1902
4.6687 | -100.2481
4.6724 | -100.303
4.6761 | -100.3596
4.6797 | -100.4074
4.6834 | -100.4501
4.6871 | -100.4942
4.6908 | -100.5375
4.6944 | -100.5725
4.6981 | -100.58
4.7018 | -100.5565
4.7054 | -100.5144
4.7091 | -100.4485
4.7128 | -100.3628
4.7165 | -100.278
4.7201 | -100.2052
4.7238 | -100.1401
4.7275 | -100.1002
4.7312 | -100.0859
4.7348 | -100.0955
4.7385 | -100.1095
4.7422 | -100.1166
4.7459 | -100.1112
4.7495 | -100.0971
4.7532 | -100.0628
4.7569 | -100.0056
4.7605 | -99.9539
4.7642 | -99.9109
4.7679 | -99.9161
4.7716 | -99.9163
4.7752 | -99.9234
4.7789 | -99.9326
4.7826 | -99.9481
4.7863 | -99.9628
4.7899 | -99.9709
4.7936 | -99.9712
4.7973 | -99.9692
4.801 | -99.9666
4.8046 | -99.9612
4.8083 | -99.9601
4.812 | -99.9679
4.8156 | -99.9895
4.8193 | -100.0129
4.823 | -100.0353
4.8267 | -100.0566
4.8303 | -100.0746
4.834 | -100.0815
4.8377 | -100.0847
4.8414 | -100.076
4.845 | -100.0573
4.8487 | -100.0289
4.8524 | -99.9992
4.8561 | -99.9632
4.8597 | -99.9229
4.8634 | -99.8813
4.8671 | -99.843
4.8707 | -99.8099
4.8744 | -99.7827
4.8781 | -99.7695
4.8818 | -99.7753
4.8854 | -99.7895
4.8891 | -99.8168
4.8928 | -99.857
4.8965 | -99.9077
4.9001 | -99.961
4.9038 | -100.0156
4.9075 | -100.0635
4.9112 | -100.1014
4.9148 | -100.1355
4.9185 | -100.1674
4.9222 | -100.1994
4.9258 | -100.2364
4.9295 | -100.2875
4.9332 | -100.352
4.9369 | -100.4304
4.9405 | -100.5207
4.9442 | -100.6068
4.9479 | -100.6807
4.9516 | -100.7428
4.9552 | -100.7878
4.9589 | -100.8183
4.9626 | -100.8438
4.9663 | -100.8632
4.9699 | -100.8744
4.9736 | -100.8802
4.9773 | -100.8843
4.9809 | -100.8826
4.9846 | -100.8652
4.9883 | -100.831
4.992 | -100.781
4.9956 | -100.7292
4.9993 | -100.676
5.003 | -100.6222
5.0067 | -100.5719
5.0103 | -100.5343
5.014 | -100.5111
5.0177 | -100.495
5.0214 | -100.4895
5.025 | -100.4899
5.0287 | -100.4912
5.0324 | -100.4817
5.036 | -100.4639
5.0397 | -100.4259
5.0434 | -100.3775
5.0471 | -100.3231
5.0507 | -100.2693
5.0544 | -100.224
5.0581 | -100.2013
5.0618 | -100.2057
5.0654 | -100.2361
5.0691 | -100.2927
5.0728 | -100.3647
5.0765 | -100.4433
5.0801 | -100.5197
5.0838 | -100.5841
5.0875 | -100.6405
5.0911 | -100.6846
5.0948 | -100.7283
5.0985 | -100.7708
5.1022 | -100.8196
5.1058 | -100.8794
5.1095 | -100.9528
5.1132 | -101.0278
5.1169 | -101.1103
5.1205 | -101.1906
5.1242 | -101.2553
5.1279 | -101.2989
5.1315 | -101.3152
5.1352 | -101.2947
5.1389 | -101.2522
5.1426 | -101.2011
5.1462 | -101.1479
5.1499 | -101.0983
5.1536 | -101.067
5.1573 | -101.055
5.1609 | -101.0629
5.1646 | -101.0858
5.1683 | -101.1206
5.172 | -101.1703
5.1756 | -101.2244
5.1793 | -101.2592
5.183 | -101.2852
5.1866 | -101.3009
5.1903 | -101.2998
5.194 | -101.2607
5.1977 | -101.1887
5.2013 | -101.0978
5.205 | -100.9777
5.2087 | -100.8371
5.2124 | -100.6842
5.216 | -100.5302
5.2197 | -100.3802
5.2234 | -100.2407
5.2271 | -100.1284
5.2307 | -100.0465
5.2344 | -100.0031
5.2381 | -99.9979
5.2417 | -100.0207
5.2454 | -100.0815
5.2491 | -100.1633
5.2528 | -100.2669
5.2564 | -100.3869
5.2601 | -100.5001
5.2638 | -100.6056
5.2675 | -100.702
5.2711 | -100.784
5.2748 | -100.833
5.2785 | -100.863
5.2822 | -100.8821
5.2858 | -100.8837
5.2895 | -100.877
5.2932 | -100.8572
5.2968 | -100.8464
5.3005 | -100.8355
5.3042 | -100.835
5.3079 | -100.837
5.3115 | -100.8454
5.3152 | -100.8611
5.3189 | -100.8848
5.3226 | -100.9155
5.3262 | -100.9359
5.3299 | -100.9498
5.3336 | -100.9524
5.3373 | -100.9504
5.3409 | -100.9461
5.3446 | -100.9378
5.3483 | -100.9299
5.3519 | -100.9275
5.3556 | -100.9313
5.3593 | -100.9323
5.363 | -100.9241
5.3666 | -100.915
5.3703 | -100.9042
5.374 | -100.8973
5.3777 | -100.888
5.3813 | -100.8811
5.385 | -100.8781
5.3887 | -100.8712
5.3924 | -100.8622
5.396 | -100.8462
5.3997 | -100.8243
5.4034 | -100.7991
5.407 | -100.7724
5.4107 | -100.7488
5.4144 | -100.7285
5.4181 | -100.7059
5.4217 | -100.6876
5.4254 | -100.6795
5.4291 | -100.6817
5.4328 | -100.6959
5.4364 | -100.7187
5.4401 | -100.7566
5.4438 | -100.7999
5.4474 | -100.8568
5.4511 | -100.9164
5.4548 | -100.9762
5.4585 | -101.0291
5.4621 | -101.0749
5.4658 | -101.1075
5.4695 | -101.1378
5.4732 | -101.1652
5.4768 | -101.1912
5.4805 | -101.2114
5.4842 | -101.2321
5.4879 | -101.2556
5.4915 | -101.2779
5.4952 | -101.3058
5.4989 | -101.3377
5.5025 | -101.3728
5.5062 | -101.3998
5.5099 | -101.4144
5.5136 | -101.4222
5.5172 | -101.4187
5.5209 | -101.4082
5.5246 | -101.3894
5.5283 | -101.3708
5.5319 | -101.3487
5.5356 | -101.3268
5.5393 | -101.3052
5.543 | -101.2948
5.5466 | -101.2824
5.5503 | -101.265
5.554 | -101.2405
5.5576 | -101.2174
5.5613 | -101.1956
5.565 | -101.1754
5.5687 | -101.145
5.5723 | -101.116
5.576 | -101.095
5.5797 | -101.0653
5.5834 | -101.0381
5.587 | -101.0221
5.5907 | -101.01
5.5944 | -100.9994
5.5981 | -100.9923
5.6017 | -100.9886
5.6054 | -100.9824
5.6091 | -100.9731
5.6127 | -100.9648
5.6164 | -100.961
5.6201 | -100.9556
5.6238 | -100.9536
5.6274 | -100.9519
5.6311 | -100.9616
5.6348 | -100.978
5.6385 | -100.9943
5.6421 | -101.0178
5.6458 | -101.0574
5.6495 | -101.1149
5.6532 | -101.1847
5.6568 | -101.2484
5.6605 | -101.3038
5.6642 | -101.3495
5.6678 | -101.376
5.6715 | -101.3743
5.6752 | -101.3529
5.6789 | -101.32
5.6825 | -101.2867
5.6862 | -101.2566
5.6899 | -101.2366
5.6936 | -101.2308
5.6972 | -101.2415
5.7009 | -101.2714
5.7046 | -101.3219
5.7083 | -101.3879
5.7119 | -101.4636
5.7156 | -101.5466
5.7193 | -101.6357
5.7229 | -101.7257
5.7266 | -101.806
5.7303 | -101.861
5.734 | -101.8885
5.7376 | -101.8976
5.7413 | -101.8813
5.745 | -101.8416
5.7487 | -101.7696
5.7523 | -101.6803
5.756 | -101.5921
5.7597 | -101.4998
5.7633 | -101.4063
5.767 | -101.3128
5.7707 | -101.2308
5.7744 | -101.1691
5.778 | -101.1256
5.7817 | -101.1003
5.7854 | -101.0965
5.7891 | -101.1181
5.7927 | -101.1565
5.7964 | -101.2047
5.8001 | -101.2474
5.8038 | -101.2921
5.8074 | -101.3224
5.8111 | -101.3368
5.8148 | -101.3376
5.8184 | -101.3254
5.8221 | -101.2982
5.8258 | -101.2513
5.8295 | -101.2223
5.8331 | -101.1961
5.8368 | -101.1778
5.8405 | -101.1602
5.8442 | -101.1449
5.8478 | -101.1377
5.8515 | -101.1392
5.8552 | -101.1527
5.8589 | -101.1705
5.8625 | -101.1877
5.8662 | -101.2077
5.8699 | -101.2312
5.8735 | -101.2594
5.8772 | -101.288
5.8809 | -101.3216
5.8846 | -101.3565
5.8882 | -101.3935
5.8919 | -101.4185
5.8956 | -101.4445
5.8993 | -101.4709
5.9029 | -101.4923
5.9066 | -101.5117
5.9103 | -101.53
5.914 | -101.543
5.9176 | -101.552
5.9213 | -101.5477
5.925 | -101.5333
5.9286 | -101.5104
5.9323 | -101.4814
5.936 | -101.4539
5.9397 | -101.4313
5.9433 | -101.4098
5.947 | -101.3961
5.9507 | -101.3965
5.9544 | -101.401
5.958 | -101.4061
5.9617 | -101.4131
5.9654 | -101.4071
5.9691 | -101.3947
5.9727 | -101.3809
5.9764 | -101.3652
5.9801 | -101.3419
5.9837 | -101.3172
5.9874 | -101.2951
5.9911 | -101.2743
5.9948 | -101.2591
5.9984 | -101.248
6.0021 | -101.2456
6.0058 | -101.245
6.0095 | -101.2426
6.0131 | -101.2441
6.0168 | -101.2481
6.0205 | -101.2546
6.0242 | -101.2573
6.0278 | -101.2561
6.0315 | -101.2545
6.0352 | -101.25
6.0388 | -101.2386
6.0425 | -101.2317
6.0462 | -101.2295
6.0499 | -101.2267
6.0535 | -101.2162
6.0572 | -101.2085
6.0609 | -101.1979
6.0646 | -101.1843
6.0682 | -101.1802
6.0719 | -101.1889
6.0756 | -101.2063
6.0793 | -101.2366
6.0829 | -101.2823
6.0866 | -101.342
6.0903 | -101.4125
6.0939 | -101.4892
6.0976 | -101.5776
6.1013 | -101.6728
6.105 | -101.7619
6.1086 | -101.8439
6.1123 | -101.9247
6.116 | -102.0057
6.1197 | -102.0797
6.1233 | -102.1506
6.127 | -102.2168
6.1307 | -102.2756
6.1344 | -102.3125
6.138 | -102.3422
6.1417 | -102.3652
6.1454 | -102.3803
6.149 | -102.379
6.1527 | -102.3689
6.1564 | -102.3436
6.1601 | -102.3062
6.1637 | -102.2653
6.1674 | -102.2247
6.1711 | -102.1855
6.1748 | -102.1503
6.1784 | -102.1248
6.1821 | -102.114
6.1858 | -102.1148
6.1894 | -102.1304
6.1931 | -102.1576
6.1968 | -102.1976
6.2005 | -102.2451
6.2041 | -102.2875
6.2078 | -102.3213
6.2115 | -102.3435
6.2152 | -102.3531
6.2188 | -102.3488
6.2225 | -102.3266
6.2262 | -102.2952
6.2299 | -102.2525
6.2335 | -102.2002
6.2372 | -102.1477
6.2409 | -102.0971
6.2445 | -102.0536
6.2482 | -102.0194
6.2519 | -101.9972
6.2556 | -101.9863
6.2592 | -101.9816
6.2629 | -101.9848
6.2666 | -101.9994
6.2703 | -102.0165
6.2739 | -102.032
6.2776 | -102.0526
6.2813 | -102.073
6.285 | -102.0849
6.2886 | -102.0901
6.2923 | -102.0875
6.296 | -102.0782
6.2996 | -102.0575
6.3033 | -102.029
6.307 | -101.9938
6.3107 | -101.9576
6.3143 | -101.9221
6.318 | -101.8837
6.3217 | -101.8455
6.3254 | -101.8142
6.329 | -101.7913
6.3327 | -101.7731
6.3364 | -101.7523
6.3401 | -101.7446
6.3437 | -101.7457
6.3474 | -101.7578
6.3511 | -101.7827
6.3547 | -101.8149
6.3584 | -101.8637
6.3621 | -101.9001
6.3658 | -101.9435
6.3694 | -101.9857
6.3731 | -102.0187
6.3768 | -102.0521
6.3805 | -102.0877
6.3841 | -102.1137
6.3878 | -102.1316
6.3915 | -102.1406
6.3952 | -102.1459
6.3988 | -102.1499
6.4025 | -102.1495
6.4062 | -102.1444
6.4098 | -102.1422
6.4135 | -102.145
6.4172 | -102.1517
6.4209 | -102.1675
6.4245 | -102.1883
6.4282 | -102.2189
6.4319 | -102.2568
6.4356 | -102.2919
6.4392 | -102.3253
6.4429 | -102.3614
6.4466 | -102.3974
6.4502 | -102.4318
6.4539 | -102.459
6.4576 | -102.4833
6.4613 | -102.5
6.4649 | -102.5114
6.4686 | -102.5206
6.4723 | -102.5245
6.476 | -102.5172
6.4796 | -102.4972
6.4833 | -102.4751
6.487 | -102.4459
6.4907 | -102.4105
6.4943 | -102.3679
6.498 | -102.3204
6.5017 | -102.2706
6.5053 | -102.2207
6.509 | -102.1717
6.5127 | -102.1283
6.5164 | -102.0895
6.52 | -102.0498
6.5237 | -102.0185
6.5274 | -101.9955
6.5311 | -101.9747
6.5347 | -101.9575
6.5384 | -101.9411
6.5421 | -101.922
6.5458 | -101.9005
6.5494 | -101.8757
6.5531 | -101.8581
6.5568 | -101.8469
6.5604 | -101.8421
6.5641 | -101.845
6.5678 | -101.8554
6.5715 | -101.8833
6.5751 | -101.9245
6.5788 | -101.9812
6.5825 | -102.0446
6.5862 | -102.1112
6.5898 | -102.1799
6.5935 | -102.2442
6.5972 | -102.3003
6.6009 | -102.3481
6.6045 | -102.3832
6.6082 | -102.3999
6.6119 | -102.4032
6.6155 | -102.3995
6.6192 | -102.3884
6.6229 | -102.3733
6.6266 | -102.3583
6.6302 | -102.3414
6.6339 | -102.3279
6.6376 | -102.3221
6.6413 | -102.3212
6.6449 | -102.3275
6.6486 | -102.3372
6.6523 | -102.3471
6.656 | -102.3598
6.6596 | -102.3774
6.6633 | -102.3937
6.667 | -102.406
6.6706 | -102.4104
6.6743 | -102.4109
6.678 | -102.4094
6.6817 | -102.4112
6.6853 | -102.4213
6.689 | -102.4383
6.6927 | -102.4652
6.6964 | -102.4967
6.7 | -102.5378
6.7037 | -102.5885
6.7074 | -102.6474
6.7111 | -102.6958
6.7147 | -102.7392
6.7184 | -102.7785
6.7221 | -102.8044
6.7257 | -102.8238
6.7294 | -102.8384
6.7331 | -102.8366
6.7368 | -102.824
6.7404 | -102.7997
6.7441 | -102.7626
6.7478 | -102.7193
6.7515 | -102.6703
6.7551 | -102.6144
6.7588 | -102.5541
6.7625 | -102.4989
6.7662 | -102.454
6.7698 | -102.4136
6.7735 | -102.3779
6.7772 | -102.3519
6.7808 | -102.3318
6.7845 | -102.3202
6.7882 | -102.316
6.7919 | -102.321
6.7955 | -102.3381
6.7992 | -102.3592
6.8029 | -102.3769
6.8066 | -102.3922
6.8102 | -102.4146
6.8139 | -102.4319
6.8176 | -102.4465
6.8213 | -102.4602
6.8249 | -102.4646
6.8286 | -102.4665
6.8323 | -102.4705
6.8359 | -102.4698
6.8396 | -102.4684
6.8433 | -102.4669
6.847 | -102.4659
6.8506 | -102.4624
6.8543 | -102.4562
6.858 | -102.4465
6.8617 | -102.4292
6.8653 | -102.4035
6.869 | -102.3746
6.8727 | -102.3451
6.8763 | -102.3103
6.88 | -102.2778
6.8837 | -102.2491
6.8874 | -102.244
6.891 | -102.2366
6.8947 | -102.2342
6.8984 | -102.2328
6.9021 | -102.239
6.9057 | -102.2571
6.9094 | -102.2804
6.9131 | -102.3049
6.9168 | -102.3265
6.9204 | -102.35
6.9241 | -102.3771
6.9278 | -102.4019
6.9314 | -102.4188
6.9351 | -102.4295
6.9388 | -102.4322
6.9425 | -102.4348
6.9461 | -102.4322
6.9498 | -102.4245
6.9535 | -102.4174
6.9572 | -102.4182
6.9608 | -102.4193
6.9645 | -102.4285
6.9682 | -102.4419
6.9719 | -102.4576
6.9755 | -102.4778
6.9792 | -102.5005
6.9829 | -102.5198
6.9865 | -102.5366
6.9902 | -102.5508
6.9939 | -102.5659
6.9976 | -102.5774
7.0012 | -102.5871
7.0049 | -102.5985
7.0086 | -102.606
7.0123 | -102.6067
7.0159 | -102.6102
7.0196 | -102.615
7.0233 | -102.6192
7.027 | -102.6205
7.0306 | -102.6203
7.0343 | -102.6257
7.038 | -102.6389
7.0416 | -102.6522
7.0453 | -102.6711
7.049 | -102.6882
7.0527 | -102.7106
7.0563 | -102.7406
7.06 | -102.7737
7.0637 | -102.8112
7.0674 | -102.8517
7.071 | -102.8855
7.0747 | -102.9132
7.0784 | -102.935
7.0821 | -102.9436
7.0857 | -102.9412
7.0894 | -102.9356
7.0931 | -102.9204
7.0967 | -102.8957
7.1004 | -102.8629
7.1041 | -102.8332
7.1078 | -102.8121
7.1114 | -102.7971
7.1151 | -102.7867
7.1188 | -102.7895
7.1225 | -102.7946
7.1261 | -102.8089
7.1298 | -102.8377
7.1335 | -102.8733
7.1372 | -102.9175
7.1408 | -102.9661
7.1445 | -103.0146
7.1482 | -103.0637
7.1518 | -103.1088
7.1555 | -103.1513
7.1592 | -103.1863
7.1629 | -103.2096
7.1665 | -103.2249
7.1702 | -103.2312
7.1739 | -103.2323
7.1776 | -103.224
7.1812 | -103.2103
7.1849 | -103.1882
7.1886 | -103.1604
7.1922 | -103.1314
7.1959 | -103.0962
7.1996 | -103.0673
7.2033 | -103.0392
7.2069 | -103.0136
7.2106 | -102.9908
7.2143 | -102.9691
7.218 | -102.9509
7.2216 | -102.9391
7.2253 | -102.9306
7.229 | -102.9261
7.2327 | -102.9274
7.2363 | -102.9328
7.24 | -102.9373
7.2437 | -102.9459
7.2473 | -102.9617
7.251 | -102.9787
7.2547 | -102.9944
7.2584 | -103.0131
7.262 | -103.0294
7.2657 | -103.0408
7.2694 | -103.0486
7.2731 | -103.0519
7.2767 | -103.0461
7.2804 | -103.0379
7.2841 | -103.024
7.2878 | -103.0048
7.2914 | -102.979
7.2951 | -102.9545
7.2988 | -102.9281
7.3024 | -102.8984
7.3061 | -102.858
7.3098 | -102.8224
7.3135 | -102.7884
7.3171 | -102.7537
7.3208 | -102.7266
7.3245 | -102.7065
7.3282 | -102.6871
7.3318 | -102.6722
7.3355 | -102.6681
7.3392 | -102.6647
7.3429 | -102.6706
7.3465 | -102.6845
7.3502 | -102.6986
7.3539 | -102.7148
7.3575 | -102.7373
7.3612 | -102.759
7.3649 | -102.7752
7.3686 | -102.7892
7.3722 | -102.8043
7.3759 | -102.8145
7.3796 | -102.818
7.3833 | -102.8178
7.3869 | -102.8091
7.3906 | -102.8007
7.3943 | -102.788
7.398 | -102.774
7.4016 | -102.7586
7.4053 | -102.7422
7.409 | -102.733
7.4126 | -102.7341
7.4163 | -102.7481
7.42 | -102.7565
7.4237 | -102.7711
7.4273 | -102.7937
7.431 | -102.8221
7.4347 | -102.8596
7.4384 | -102.9098
7.442 | -102.9618
7.4457 | -103.0179
7.4494 | -103.0824
7.4531 | -103.1465
7.4567 | -103.2141
7.4604 | -103.2834
7.4641 | -103.3505
7.4677 | -103.412
7.4714 | -103.4737
7.4751 | -103.5324
7.4788 | -103.5842
7.4824 | -103.6265
7.4861 | -103.6663
7.4898 | -103.6996
7.4935 | -103.7277
7.4971 | -103.7534
7.5008 | -103.771
7.5045 | -103.7821
7.5081 | -103.7922
7.5118 | -103.798
7.5155 | -103.7977
7.5192 | -103.7962
7.5228 | -103.786
7.5265 | -103.7743
7.5302 | -103.7509
7.5339 | -103.7226
7.5375 | -103.6893
7.5412 | -103.6497
7.5449 | -103.6016
7.5486 | -103.5494
7.5522 | -103.4966
7.5559 | -103.4454
7.5596 | -103.3877
7.5632 | -103.3354
7.5669 | -103.2878
7.5706 | -103.2421
7.5743 | -103.204
7.5779 | -103.1706
7.5816 | -103.1432
7.5853 | -103.1197
7.589 | -103.1026
7.5926 | -103.0991
7.5963 | -103.1036
7.6 | -103.1148
7.6037 | -103.1342
7.6073 | -103.1547
7.611 | -103.1762
7.6147 | -103.2052
7.6183 | -103.2374
7.622 | -103.2658
7.6257 | -103.295
7.6294 | -103.3266
7.633 | -103.3423
7.6367 | -103.3577
7.6404 | -103.3679
7.6441 | -103.3717
7.6477 | -103.3687
7.6514 | -103.3629
7.6551 | -103.3541
7.6588 | -103.3462
7.6624 | -103.3379
7.6661 | -103.33
7.6698 | -103.3301
7.6734 | -103.3361
7.6771 | -103.3429
7.6808 | -103.3505
7.6845 | -103.3681
7.6881 | -103.3925
7.6918 | -103.4269
7.6955 | -103.4641
7.6992 | -103.5023
7.7028 | -103.5425
7.7065 | -103.5743
7.7102 | -103.6045
7.7139 | -103.6288
7.7175 | -103.6461
7.7212 | -103.6575
7.7249 | -103.6587
7.7285 | -103.6402
7.7322 | -103.611
7.7359 | -103.5786
7.7396 | -103.5374
7.7432 | -103.4944
7.7469 | -103.4572
7.7506 | -103.4178
7.7543 | -103.3777
7.7579 | -103.3421
7.7616 | -103.3054
7.7653 | -103.2741
7.769 | -103.2559
7.7726 | -103.246
7.7763 | -103.2427
7.78 | -103.2421
7.7836 | -103.2448
7.7873 | -103.2584
7.791 | -103.2788
7.7947 | -103.3068
7.7983 | -103.339
7.802 | -103.3742
7.8057 | -103.4061
7.8094 | -103.4406
7.813 | -103.4796
7.8167 | -103.518
7.8204 | -103.5554
7.8241 | -103.587
7.8277 | -103.6148
7.8314 | -103.6342
7.8351 | -103.6462
7.8387 | -103.6581
7.8424 | -103.6694
7.8461 | -103.6745
7.8498 | -103.6658
7.8534 | -103.6562
7.8571 | -103.6458
7.8608 | -103.6355
7.8645 | -103.6276
7.8681 | -103.6238
7.8718 | -103.6192
7.8755 | -103.6145
7.8792 | -103.6083
7.8828 | -103.6016
7.8865 | -103.596
7.8902 | -103.5949
7.8938 | -103.6024
7.8975 | -103.6076
7.9012 | -103.6092
7.9049 | -103.6084
7.9085 | -103.6072
7.9122 | -103.6054
7.9159 | -103.6031
7.9196 | -103.6023
7.9232 | -103.5997
7.9269 | -103.5882
7.9306 | -103.5769
7.9342 | -103.5646
7.9379 | -103.5498
7.9416 | -103.53
7.9453 | -103.5159
7.9489 | -103.5076
7.9526 | -103.4966
7.9563 | -103.4868
7.96 | -103.4818
7.9636 | -103.4851
7.9673 | -103.4944
7.971 | -103.5011
7.9747 | -103.5108
7.9783 | -103.5211
7.982 | -103.5326
7.9857 | -103.5464
7.9893 | -103.5633
7.993 | -103.5817
7.9967 | -103.5959
8.0004 | -103.6056
8.004 | -103.6182
8.0077 | -103.6261
8.0114 | -103.6298
8.0151 | -103.632
8.0187 | -103.6356
8.0224 | -103.6393
8.0261 | -103.6388
8.0298 | -103.6305
8.0334 | -103.6189
8.0371 | -103.607
8.0408 | -103.596
8.0444 | -103.5841
8.0481 | -103.5627
8.0518 | -103.5412
8.0555 | -103.5225
8.0591 | -103.4997
8.0628 | -103.4689
8.0665 | -103.4363
8.0702 | -103.4063
8.0738 | -103.3776
8.0775 | -103.3399
8.0812 | -103.3047
8.0849 | -103.2708
8.0885 | -103.2418
8.0922 | -103.2206
8.0959 | -103.2015
8.0995 | -103.1867
8.1032 | -103.1787
8.1069 | -103.1758
8.1106 | -103.1765
8.1142 | -103.1817
8.1179 | -103.1963
8.1216 | -103.218
8.1253 | -103.2378
8.1289 | -103.2625
8.1326 | -103.2898
8.1363 | -103.322
8.14 | -103.3512
8.1436 | -103.3823
8.1473 | -103.4151
8.151 | -103.45
8.1546 | -103.4827
8.1583 | -103.5175
8.162 | -103.5467
8.1657 | -103.5745
8.1693 | -103.6066
8.173 | -103.6395
8.1767 | -103.6712
8.1804 | -103.7069
8.184 | -103.7459
8.1877 | -103.7855
8.1914 | -103.8198
8.1951 | -103.8537
8.1987 | -103.8905
8.2024 | -103.9157
8.2061 | -103.9364
8.2097 | -103.9551
8.2134 | -103.9613
8.2171 | -103.9588
8.2208 | -103.9503
8.2244 | -103.9329
8.2281 | -103.9108
8.2318 | -103.8844
8.2355 | -103.8498
8.2391 | -103.8035
8.2428 | -103.7523
8.2465 | -103.7047
8.2501 | -103.657
8.2538 | -103.6075
8.2575 | -103.5605
8.2612 | -103.52
8.2648 | -103.4825
8.2685 | -103.4507
8.2722 | -103.4263
8.2759 | -103.4123
8.2795 | -103.406
8.2832 | -103.4092
8.2869 | -103.4203
8.2906 | -103.4376
8.2942 | -103.461
8.2979 | -103.4912
8.3016 | -103.5258
8.3052 | -103.5589
8.3089 | -103.5939
8.3126 | -103.6324
8.3163 | -103.6709
8.3199 | -103.7076
8.3236 | -103.7416
8.3273 | -103.7647
8.331 | -103.7776
8.3346 | -103.7887
8.3383 | -103.796
8.342 | -103.7986
8.3457 | -103.7929
8.3493 | -103.7778
8.353 | -103.7623
8.3567 | -103.7498
8.3603 | -103.7426
8.364 | -103.7389
8.3677 | -103.735
8.3714 | -103.7348
8.375 | -103.7421
8.3787 | -103.7567
8.3824 | -103.7773
8.3861 | -103.8032
8.3897 | -103.8327
8.3934 | -103.8705
8.3971 | -103.9117
8.4008 | -103.9522
8.4044 | -103.9978
8.4081 | -104.0424
8.4118 | -104.0788
8.4154 | -104.1103
8.4191 | -104.1354
8.4228 | -104.1581
8.4265 | -104.1753
8.4301 | -104.1802
8.4338 | -104.1722
8.4375 | -104.1548
8.4412 | -104.1309
8.4448 | -104.1008
8.4485 | -104.0656
8.4522 | -104.0277
8.4559 | -103.9876
8.4595 | -103.9429
8.4632 | -103.8997
8.4669 | -103.8605
8.4705 | -103.8298
8.4742 | -103.8006
8.4779 | -103.7857
8.4816 | -103.7598
8.4852 | -103.7408
8.4889 | -103.7256
8.4926 | -103.7163
8.4963 | -103.7121
8.4999 | -103.7085
8.5036 | -103.7042
8.5073 | -103.7002
8.511 | -103.6983
8.5146 | -103.7011
8.5183 | -103.7024
8.522 | -103.7042
8.5256 | -103.7113
8.5293 | -103.7174
8.533 | -103.7234
8.5367 | -103.7355
8.5403 | -103.7494
8.544 | -103.7666
8.5477 | -103.7849
8.5514 | -103.8088
8.555 | -103.8353
8.5587 | -103.8628
8.5624 | -103.8883
8.5661 | -103.916
8.5697 | -103.9446
8.5734 | -103.9667
8.5771 | -103.9839
8.5807 | -104.0003
8.5844 | -104.0166
8.5881 | -104.0318
8.5918 | -104.0453
8.5954 | -104.0541
8.5991 | -104.0619
8.6028 | -104.0663
8.6065 | -104.0729
8.6101 | -104.0742
8.6138 | -104.0708
8.6175 | -104.0712
8.6211 | -104.0687
8.6248 | -104.0664
8.6285 | -104.0688
8.6322 | -104.0685
8.6358 | -104.0681
8.6395 | -104.0634
8.6432 | -104.0532
8.6469 | -104.0404
8.6505 | -104.0263
8.6542 | -104.0068
8.6579 | -103.9859
8.6616 | -103.9701
8.6652 | -103.9484
8.6689 | -103.9221
8.6726 | -103.8937
8.6762 | -103.8685
8.6799 | -103.8441
8.6836 | -103.822
8.6873 | -103.8071
8.6909 | -103.7986
8.6946 | -103.7957
8.6983 | -103.7973
8.702 | -103.8103
8.7056 | -103.8272
8.7093 | -103.8443
8.713 | -103.8677
8.7167 | -103.889
8.7203 | -103.9145
8.724 | -103.9452
8.7277 | -103.9777
8.7313 | -104.0094
8.735 | -104.0424
8.7387 | -104.0742
8.7424 | -104.0981
8.746 | -104.1154
8.7497 | -104.1307
8.7534 | -104.1408
8.7571 | -104.1451
8.7607 | -104.1441
8.7644 | -104.1381
8.7681 | -104.1335
8.7718 | -104.1216
8.7754 | -104.1034
8.7791 | -104.0855
8.7828 | -104.0642
8.7864 | -104.0429
8.7901 | -104.0181
8.7938 | -103.9914
8.7975 | -103.9689
8.8011 | -103.9499
8.8048 | -103.9278
8.8085 | -103.9081
8.8122 | -103.894
8.8158 | -103.8808
8.8195 | -103.8676
8.8232 | -103.8535
8.8269 | -103.8456
8.8305 | -103.8413
8.8342 | -103.8404
8.8379 | -103.8454
8.8415 | -103.8491
8.8452 | -103.8636
8.8489 | -103.8819
8.8526 | -103.9023
8.8562 | -103.9264
8.8599 | -103.9521
8.8636 | -103.979
8.8673 | -104.0081
8.8709 | -104.0371
8.8746 | -104.0621
8.8783 | -104.0837
8.882 | -104.1059
8.8856 | -104.1242
8.8893 | -104.1402
8.893 | -104.1527
8.8966 | -104.1647
8.9003 | -104.1802
8.904 | -104.1912
8.9077 | -104.1977
8.9113 | -104.2111
8.915 | -104.2229
8.9187 | -104.2308
8.9224 | -104.2386
8.926 | -104.2425
8.9297 | -104.2387
8.9334 | -104.2344
8.937 | -104.2248
8.9407 | -104.209
8.9444 | -104.1939
8.9481 | -104.1812
8.9517 | -104.1689
8.9554 | -104.1547
8.9591 | -104.1397
8.9628 | -104.1199
8.9664 | -104.1001
8.9701 | -104.0828
8.9738 | -104.058
8.9775 | -104.0352
8.9811 | -104.0137
8.9848 | -103.9947
8.9885 | -103.9825
8.9921 | -103.9712
8.9958 | -103.9661
8.9995 | -103.9671
9.0032 | -103.9743
9.0068 | -103.995
9.0105 | -104.0074
9.0142 | -104.0172
9.0179 | -104.0326
9.0215 | -104.051
9.0252 | -104.0712
9.0289 | -104.0873
9.0326 | -104.0973
9.0362 | -104.1123
9.0399 | -104.126
9.0436 | -104.1337
9.0472 | -104.1391
9.0509 | -104.1471
9.0546 | -104.1536
9.0583 | -104.155
9.0619 | -104.1605
9.0656 | -104.161
9.0693 | -104.1563
9.073 | -104.154
9.0766 | -104.1519
9.0803 | -104.1434
9.084 | -104.1372
9.0877 | -104.1334
9.0913 | -104.1314
9.095 | -104.134
9.0987 | -104.1338
9.1023 | -104.135
9.106 | -104.1402
9.1097 | -104.1435
9.1134 | -104.1504
9.117 | -104.1583
9.1207 | -104.1697
9.1244 | -104.1832
9.1281 | -104.194
9.1317 | -104.209
9.1354 | -104.2257
9.1391 | -104.2467
9.1428 | -104.2699
9.1464 | -104.2993
9.1501 | -104.327
9.1538 | -104.3497
9.1574 | -104.3746
9.1611 | -104.4008
9.1648 | -104.4247
9.1685 | -104.4506
9.1721 | -104.4674
9.1758 | -104.477
9.1795 | -104.4765
9.1832 | -104.4707
9.1868 | -104.4598
9.1905 | -104.4474
9.1942 | -104.428
9.1979 | -104.4037
9.2015 | -104.3824
9.2052 | -104.3631
9.2089 | -104.3402
9.2125 | -104.3145
9.2162 | -104.2965
9.2199 | -104.2832
9.2236 | -104.2702
9.2272 | -104.2627
9.2309 | -104.2589
9.2346 | -104.2605
9.2383 | -104.2652
9.2419 | -104.2728
9.2456 | -104.2811
9.2493 | -104.2897
9.253 | -104.303
9.2566 | -104.3184
9.2603 | -104.334
9.264 | -104.351
9.2676 | -104.3648
9.2713 | -104.3734
9.275 | -104.3796
9.2787 | -104.3827
9.2823 | -104.381
9.286 | -104.3809
9.2897 | -104.3832
9.2934 | -104.3818
9.297 | -104.3799
9.3007 | -104.3794
9.3044 | -104.3778
9.308 | -104.3733
9.3117 | -104.367
9.3154 | -104.3712
9.3191 | -104.3761
9.3227 | -104.3736
9.3264 | -104.3669
9.3301 | -104.3684
9.3338 | -104.3677
9.3374 | -104.3645
9.3411 | -104.3609
9.3448 | -104.3545
9.3485 | -104.3491
9.3521 | -104.3406
9.3558 | -104.3335
9.3595 | -104.3284
9.3631 | -104.319
9.3668 | -104.3118
9.3705 | -104.3024
9.3742 | -104.2912
9.3778 | -104.2816
9.3815 | -104.2749
9.3852 | -104.2664
9.3889 | -104.2545
9.3925 | -104.2447
9.3962 | -104.2402
9.3999 | -104.2384
9.4036 | -104.2381
9.4072 | -104.2399
9.4109 | -104.2435
9.4146 | -104.2506
9.4182 | -104.262
9.4219 | -104.2757
9.4256 | -104.2876
9.4293 | -104.3051
9.4329 | -104.3237
9.4366 | -104.343
9.4403 | -104.364
9.444 | -104.3868
9.4476 | -104.4056
9.4513 | -104.4186
9.455 | -104.4304
9.4587 | -104.4453
9.4623 | -104.4599
9.466 | -104.4662
9.4697 | -104.4682
9.4733 | -104.4738
9.477 | -104.4777
9.4807 | -104.4812
9.4844 | -104.4837
9.488 | -104.484
9.4917 | -104.4768
9.4954 | -104.4705
9.4991 | -104.4697
9.5027 | -104.4696
9.5064 | -104.4681
9.5101 | -104.4673
9.5138 | -104.4681
9.5174 | -104.4667
9.5211 | -104.4641
9.5248 | -104.4604
9.5284 | -104.455
9.5321 | -104.4514
9.5358 | -104.44
9.5395 | -104.4336
9.5431 | -104.4248
9.5468 | -104.4178
9.5505 | -104.4084
9.5542 | -104.3977
9.5578 | -104.3857
9.5615 | -104.375
9.5652 | -104.3676
9.5689 | -104.3596
9.5725 | -104.354
9.5762 | -104.3552
9.5799 | -104.3607
9.5835 | -104.3738
9.5872 | -104.3976
9.5909 | -104.4206
9.5946 | -104.4447
9.5982 | -104.4742
9.6019 | -104.5051
9.6056 | -104.5358
9.6093 | -104.5709
9.6129 | -104.5999
9.6166 | -104.6275
9.6203 | -104.6528
9.624 | -104.6724
9.6276 | -104.6893
9.6313 | -104.7015
9.635 | -104.7097
9.6386 | -104.7108
9.6423 | -104.7053
9.646 | -104.699
9.6497 | -104.6868
9.6533 | -104.6676
9.657 | -104.6485
9.6607 | -104.6292
9.6644 | -104.607
9.668 | -104.5802
9.6717 | -104.5501
9.6754 | -104.5221
9.679 | -104.4975
9.6827 | -104.4747
9.6864 | -104.4568
9.6901 | -104.4439
9.6937 | -104.4355
9.6974 | -104.431
9.7011 | -104.4338
9.7048 | -104.4433
9.7084 | -104.4575
9.7121 | -104.4788
9.7158 | -104.5053
9.7195 | -104.5339
9.7231 | -104.5663
9.7268 | -104.5954
9.7305 | -104.6182
9.7341 | -104.6482
9.7378 | -104.6802
9.7415 | -104.7109
9.7452 | -104.7389
9.7488 | -104.7666
9.7525 | -104.7902
9.7562 | -104.8134
9.7599 | -104.835
9.7635 | -104.8527
9.7672 | -104.8692
9.7709 | -104.885
9.7746 | -104.8979
9.7782 | -104.9035
9.7819 | -104.9087
9.7856 | -104.9127
9.7892 | -104.9117
9.7929 | -104.9142
9.7966 | -104.9151
9.8003 | -104.9151
9.8039 | -104.9141
9.8076 | -104.9084
9.8113 | -104.8958
9.815 | -104.8832
9.8186 | -104.8697
9.8223 | -104.8529
9.826 | -104.8324
9.8297 | -104.8068
9.8333 | -104.7761
9.837 | -104.7441
9.8407 | -104.7123
9.8443 | -104.6798
9.848 | -104.6471
9.8517 | -104.6149
9.8554 | -104.5847
9.859 | -104.5559
9.8627 | -104.5343
9.8664 | -104.5168
9.8701 | -104.5044
9.8737 | -104.5025
9.8774 | -104.4996
9.8811 | -104.5037
9.8848 | -104.5137
9.8884 | -104.5287
9.8921 | -104.5393
9.8958 | -104.548
9.8994 | -104.5597
9.9031 | -104.5746
9.9068 | -104.5889
9.9105 | -104.6051
9.9141 | -104.6171
9.9178 | -104.6287
9.9215 | -104.6394
9.9252 | -104.6424
9.9288 | -104.6426
9.9325 | -104.6407
9.9362 | -104.6348
9.9398 | -104.6277
9.9435 | -104.6155
9.9472 | -104.6015
9.9509 | -104.5894
9.9545 | -104.5787
9.9582 | -104.5622
9.9619 | -104.5411
9.9656 | -104.5225
9.9692 | -104.5055
9.9729 | -104.4904
9.9766 | -104.474
9.9803 | -104.4573
9.9839 | -104.4397
9.9876 | -104.4234
9.9913 | -104.41
9.9949 | -104.3985
9.9986 | -104.3874
10.0023 | -104.3801
10.006 | -104.3745
10.0096 | -104.3731
10.0133 | -104.3775
10.017 | -104.3828
10.0207 | -104.3936
10.0243 | -104.406
10.028 | -104.4236
10.0317 | -104.4467
10.0354 | -104.4777
10.039 | -104.5143
10.0427 | -104.5527
10.0464 | -104.5921
10.05 | -104.6352
10.0537 | -104.6756
10.0574 | -104.7212
10.0611 | -104.7656
10.0647 | -104.8066
10.0684 | -104.8363
10.0721 | -104.8746
10.0758 | -104.9079
10.0794 | -104.9335
10.0831 | -104.9506
10.0868 | -104.9609
10.0905 | -104.9654
10.0941 | -104.966
10.0978 | -104.9605
10.1015 | -104.9488
10.1051 | -104.9339
10.1088 | -104.9123
10.1125 | -104.8832
10.1162 | -104.8548
10.1198 | -104.8237
10.1235 | -104.7889
10.1272 | -104.7521
10.1309 | -104.7152
10.1345 | -104.6785
10.1382 | -104.6446
10.1419 | -104.6175
10.1456 | -104.5965
10.1492 | -104.5761
10.1529 | -104.5569
10.1566 | -104.5405
10.1602 | -104.5294
10.1639 | -104.5258
10.1676 | -104.5272
10.1713 | -104.5338
10.1749 | -104.5443
10.1786 | -104.5585
10.1823 | -104.577
10.186 | -104.5946
10.1896 | -104.6156
10.1933 | -104.6399
10.197 | -104.6652
10.2007 | -104.6945
10.2043 | -104.7222
10.208 | -104.7497
10.2117 | -104.7783
10.2153 | -104.8053
10.219 | -104.8326
10.2227 | -104.8568
10.2264 | -104.8807
10.23 | -104.9045
10.2337 | -104.9218
10.2374 | -104.9394
10.2411 | -104.9569
10.2447 | -104.9663
10.2484 | -104.969
10.2521 | -104.971
10.2558 | -104.9716
10.2594 | -104.9755
10.2631 | -104.9764
10.2668 | -104.9689
10.2704 | -104.9646
10.2741 | -104.9608
10.2778 | -104.9575
10.2815 | -104.9523
10.2851 | -104.9431
10.2888 | -104.9303
10.2925 | -104.9141
10.2962 | -104.8952
10.2998 | -104.8803
10.3035 | -104.8679
10.3072 | -104.8551
10.3109 | -104.839
10.3145 | -104.8232
10.3182 | -104.8059
10.3219 | -104.7936
10.3255 | -104.7855
10.3292 | -104.7773
10.3329 | -104.7691
10.3366 | -104.7634
10.3402 | -104.758
10.3439 | -104.7497
10.3476 | -104.744
10.3513 | -104.7391
10.3549 | -104.7347
10.3586 | -104.7307
10.3623 | -104.7279
10.3659 | -104.7275
10.3696 | -104.7263
10.3733 | -104.7268
10.377 | -104.7264
10.3806 | -104.7237
10.3843 | -104.7229
10.388 | -104.7215
10.3917 | -104.7188
10.3953 | -104.7207
10.399 | -104.7255
10.4027 | -104.7277
10.4064 | -104.7288
10.41 | -104.7298
10.4137 | -104.7342
10.4174 | -104.7435
10.421 | -104.7519
10.4247 | -104.7627
10.4284 | -104.7708
10.4321 | -104.784
10.4357 | -104.8001
10.4394 | -104.8184
10.4431 | -104.8402
10.4468 | -104.8652
10.4504 | -104.8857
10.4541 | -104.9054
10.4578 | -104.9256
10.4615 | -104.946
10.4651 | -104.9636
10.4688 | -104.9826
10.4725 | -104.9959
10.4761 | -105.0042
10.4798 | -105.0112
10.4835 | -105.014
10.4872 | -105.015
10.4908 | -105.0109
10.4945 | -105.0076
10.4982 | -104.9993
10.5019 | -104.9831
10.5055 | -104.9633
10.5092 | -104.9446
10.5129 | -104.9253
10.5166 | -104.9048
10.5202 | -104.8886
10.5239 | -104.8733
10.5276 | -104.8581
10.5312 | -104.8475
10.5349 | -104.8359
10.5386 | -104.8239
10.5423 | -104.8191
10.5459 | -104.8167
10.5496 | -104.8162
10.5533 | -104.8206
10.557 | -104.8277
10.5606 | -104.8364
10.5643 | -104.8458
10.568 | -104.8615
10.5717 | -104.8793
10.5753 | -104.9019
10.579 | -104.9277
10.5827 | -104.9503
10.5863 | -104.974
10.59 | -104.9988
10.5937 | -105.022
10.5974 | -105.0425
10.601 | -105.062
10.6047 | -105.0859
10.6084 | -105.1046
10.6121 | -105.1199
10.6157 | -105.1343
10.6194 | -105.1473
10.6231 | -105.1576
10.6268 | -105.164
10.6304 | -105.1677
10.6341 | -105.1698
10.6378 | -105.1697
10.6414 | -105.1636
10.6451 | -105.1575
10.6488 | -105.1486
10.6525 | -105.1375
10.6561 | -105.1264
10.6598 | -105.1147
10.6635 | -105.0959
10.6672 | -105.0742
10.6708 | -105.0532
10.6745 | -105.0326
10.6782 | -105.0158
10.6818 | -105.0012
10.6855 | -104.9848
10.6892 | -104.9725
10.6929 | -104.9627
10.6965 | -104.9517
10.7002 | -104.9416
10.7039 | -104.9369
10.7076 | -104.9336
10.7112 | -104.9316
10.7149 | -104.9348
10.7186 | -104.9435
10.7223 | -104.9502
10.7259 | -104.9627
10.7296 | -104.9762
10.7333 | -104.9903
10.7369 | -105.005
10.7406 | -105.0208
10.7443 | -105.0384
10.748 | -105.0548
10.7516 | -105.07
10.7553 | -105.087
10.759 | -105.0993
10.7627 | -105.1067
10.7663 | -105.1126
10.77 | -105.1147
10.7737 | -105.115
10.7774 | -105.1107
10.781 | -105.102
10.7847 | -105.0927
10.7884 | -105.0797
10.792 | -105.0615
10.7957 | -105.0427
10.7994 | -105.023
10.8031 | -104.9985
10.8067 | -104.9728
10.8104 | -104.9477
10.8141 | -104.9238
10.8178 | -104.8997
10.8214 | -104.8818
10.8251 | -104.8663
10.8288 | -104.8516
10.8325 | -104.8384
10.8361 | -104.8294
10.8398 | -104.8226
10.8435 | -104.8193
10.8471 | -104.8202
10.8508 | -104.8229
10.8545 | -104.8261
10.8582 | -104.8327
10.8618 | -104.8403
10.8655 | -104.8492
10.8692 | -104.8651
10.8729 | -104.8817
10.8765 | -104.8989
10.8802 | -104.9168
10.8839 | -104.934
10.8876 | -104.9524
10.8912 | -104.9727
10.8949 | -104.9932
10.8986 | -105.0104
10.9022 | -105.0234
10.9059 | -105.0332
10.9096 | -105.0413
10.9133 | -105.0445
10.9169 | -105.0475
10.9206 | -105.0449
10.9243 | -105.0395
10.928 | -105.0348
10.9316 | -105.0265
10.9353 | -105.0192
10.939 | -105.0116
10.9427 | -105.0025
10.9463 | -104.9953
10.95 | -104.992
10.9537 | -104.989
10.9573 | -104.9882
10.961 | -104.991
10.9647 | -104.9962
10.9684 | -105.0022
10.972 | -105.0099
10.9757 | -105.0231
10.9794 | -105.0355
10.9831 | -105.0491
10.9867 | -105.0674
10.9904 | -105.0899
10.9941 | -105.1156
10.9977 | -105.1478
11.0014 | -105.1759
11.0051 | -105.2021
11.0088 | -105.2293
11.0124 | -105.2555
11.0161 | -105.278
11.0198 | -105.2968
11.0235 | -105.3091
11.0271 | -105.3116
11.0308 | -105.3095
11.0345 | -105.3044
11.0382 | -105.2959
11.0418 | -105.2812
11.0455 | -105.2641
11.0492 | -105.2446
11.0528 | -105.2226
11.0565 | -105.2012
11.0602 | -105.1804
11.0639 | -105.1608
11.0675 | -105.1404
11.0712 | -105.1243
11.0749 | -105.1127
11.0786 | -105.1054
11.0822 | -105.1
11.0859 | -105.0982
11.0896 | -105.0986
11.0933 | -105.0993
11.0969 | -105.1073
11.1006 | -105.1223
11.1043 | -105.1403
11.1079 | -105.1603
11.1116 | -105.1824
11.1153 | -105.2004
11.119 | -105.222
11.1226 | -105.2455
11.1263 | -105.2514
11.13 | -105.2644
11.1337 | -105.2737
11.1373 | -105.2776
11.141 | -105.2793
11.1447 | -105.2781
11.1484 | -105.2747
11.152 | -105.2673
11.1557 | -105.2584
11.1594 | -105.2421
11.163 | -105.2288
11.1667 | -105.2077
11.1704 | -105.1846
11.1741 | -105.1616
11.1777 | -105.1375
11.1814 | -105.1089
11.1851 | -105.0825
11.1888 | -105.0572
11.1924 | -105.0315
11.1961 | -105.0069
11.1998 | -104.9831
11.2035 | -104.9598
11.2071 | -104.9382
11.2108 | -104.9199
11.2145 | -104.9052
11.2181 | -104.8934
11.2218 | -104.8847
11.2255 | -104.8789
11.2292 | -104.8756
11.2328 | -104.8762
11.2365 | -104.8784
11.2402 | -104.8871
11.2439 | -104.8983
11.2475 | -104.9126
11.2512 | -104.9318
11.2549 | -104.9569
11.2586 | -104.9837
11.2622 | -105.0108
11.2659 | -105.0389
11.2696 | -105.07
11.2732 | -105.1021
11.2769 | -105.1396
11.2806 | -105.175
11.2843 | -105.2092
11.2879 | -105.2454
11.2916 | -105.282
11.2953 | -105.3169
11.299 | -105.3476
11.3026 | -105.3809
11.3063 | -105.4103
11.31 | -105.4354
11.3137 | -105.4544
11.3173 | -105.4679
11.321 | -105.4739
11.3247 | -105.4839
11.3283 | -105.4896
11.332 | -105.4953
11.3357 | -105.4976
11.3394 | -105.4923
11.343 | -105.4868
11.3467 | -105.4807
11.3504 | -105.4669
11.3541 | -105.4497
11.3577 | -105.4336
11.3614 | -105.4175
11.3651 | -105.4046
11.3688 | -105.397
11.3724 | -105.3888
11.3761 | -105.3822
11.3798 | -105.3771
11.3834 | -105.3717
11.3871 | -105.3692
11.3908 | -105.3733
11.3945 | -105.3815
11.3981 | -105.3897
11.4018 | -105.3994
11.4055 | -105.4063
11.4092 | -105.4126
11.4128 | -105.4183
11.4165 | -105.4232
11.4202 | -105.4281
11.4238 | -105.4345
11.4275 | -105.439
11.4312 | -105.4437
11.4349 | -105.4468
11.4385 | -105.4504
11.4422 | -105.4508
11.4459 | -105.4528
11.4496 | -105.4529
11.4532 | -105.448
11.4569 | -105.4435
11.4606 | -105.4382
11.4643 | -105.4312
11.4679 | -105.4187
11.4716 | -105.409
11.4753 | -105.4028
11.4789 | -105.3988
11.4826 | -105.3975
11.4863 | -105.3957
11.49 | -105.3904
11.4936 | -105.3865
11.4973 | -105.3853
11.501 | -105.3817
11.5047 | -105.3832
11.5083 | -105.3848
11.512 | -105.3846
11.5157 | -105.3851
11.5194 | -105.3873
11.523 | -105.3869
11.5267 | -105.3869
11.5304 | -105.3862
11.534 | -105.3858
11.5377 | -105.3861
11.5414 | -105.3856
11.5451 | -105.3852
11.5487 | -105.388
11.5524 | -105.3886
11.5561 | -105.3867
11.5598 | -105.3867
11.5634 | -105.3863
11.5671 | -105.3851
11.5708 | -105.3872
11.5745 | -105.3937
11.5781 | -105.3962
11.5818 | -105.398
11.5855 | -105.3989
11.5891 | -105.3923
11.5928 | -105.3854
11.5965 | -105.3806
11.6002 | -105.3731
11.6038 | -105.3628
11.6075 | -105.3555
11.6112 | -105.3404
11.6149 | -105.3243
11.6185 | -105.3113
11.6222 | -105.2958
11.6259 | -105.2798
11.6296 | -105.2654
11.6332 | -105.2488
11.6369 | -105.2329
11.6406 | -105.2178
11.6442 | -105.203
11.6479 | -105.188
11.6516 | -105.1778
11.6553 | -105.1802
11.6589 | -105.1762
11.6626 | -105.1744
11.6663 | -105.1752
11.67 | -105.1826
11.6736 | -105.1868
11.6773 | -105.1913
11.681 | -105.2028
11.6846 | -105.2189
11.6883 | -105.234
11.692 | -105.2465
11.6957 | -105.2571
11.6993 | -105.269
11.703 | -105.2815
11.7067 | -105.2933
11.7104 | -105.3043
11.714 | -105.3137
11.7177 | -105.3226
11.7214 | -105.331
11.7251 | -105.3363
11.7287 | -105.3366
11.7324 | -105.337
11.7361 | -105.3348
11.7397 | -105.3316
11.7434 | -105.3251
11.7471 | -105.3196
11.7508 | -105.3158
11.7544 | -105.3092
11.7581 | -105.3046
11.7618 | -105.2977
11.7655 | -105.2886
11.7691 | -105.2774
11.7728 | -105.2667
11.7765 | -105.2545
11.7802 | -105.2449
11.7838 | -105.2342
11.7875 | -105.2198
11.7912 | -105.2024
11.7948 | -105.1863
11.7985 | -105.1708
11.8022 | -105.1537
11.8059 | -105.1343
11.8095 | -105.115
11.8132 | -105.0987
11.8169 | -105.0856
11.8206 | -105.0696
11.8242 | -105.0529
11.8279 | -105.0388
11.8316 | -105.0213
11.8353 | -105.0068
11.8389 | -104.9918
11.8426 | -104.9777
11.8463 | -104.9714
11.8499 | -104.9659
11.8536 | -104.9593
11.8573 | -104.953
11.861 | -104.9494
11.8646 | -104.952
11.8683 | -104.9554
11.872 | -104.9581
11.8757 | -104.9641
11.8793 | -104.9732
11.883 | -104.9852
11.8867 | -105.0016
11.8904 | -105.0201
11.894 | -105.0378
11.8977 | -105.0579
11.9014 | -105.0832
11.905 | -105.1098
11.9087 | -105.1377
11.9124 | -105.1621
11.9161 | -105.1887
11.9197 | -105.22
11.9234 | -105.2493
11.9271 | -105.2801
11.9308 | -105.3097
11.9344 | -105.3381
11.9381 | -105.3674
11.9418 | -105.396
11.9455 | -105.422
11.9491 | -105.4421
11.9528 | -105.4589
11.9565 | -105.4727
11.9601 | -105.4863
11.9638 | -105.4973
11.9675 | -105.5019
11.9712 | -105.5042
11.9748 | -105.5055
11.9785 | -105.5049
11.9822 | -105.4996
11.9859 | -105.4917
11.9895 | -105.4841
11.9932 | -105.4774
11.9969 | -105.4685
12.0006 | -105.4552
12.0042 | -105.4468
12.0079 | -105.4379
12.0116 | -105.4305
12.0152 | -105.4223
12.0189 | -105.4151
12.0226 | -105.4128
12.0263 | -105.4113
12.0299 | -105.4108
12.0336 | -105.4105
12.0373 | -105.411
12.041 | -105.4135
12.0446 | -105.4189
12.0483 | -105.4209
12.052 | -105.4256
12.0556 | -105.4302
12.0593 | -105.4329
12.063 | -105.4358
12.0667 | -105.4432
12.0703 | -105.45
12.074 | -105.4561
12.0777 | -105.4627
12.0814 | -105.4726
12.085 | -105.4831
12.0887 | -105.488
12.0924 | -105.4916
12.0961 | -105.4969
12.0997 | -105.5015
12.1034 | -105.5113
12.1071 | -105.5244
12.1107 | -105.5325
12.1144 | -105.5438
12.1181 | -105.5569
12.1218 | -105.5729
12.1254 | -105.5906
12.1291 | -105.6071
12.1328 | -105.6198
12.1365 | -105.6324
12.1401 | -105.6478
12.1438 | -105.6617
12.1475 | -105.6726
12.1512 | -105.6851
12.1548 | -105.6976
12.1585 | -105.7109
12.1622 | -105.7212
12.1658 | -105.7334
12.1695 | -105.7465
12.1732 | -105.7626
12.1769 | -105.7781
12.1805 | -105.7922
12.1842 | -105.8028
12.1879 | -105.8128
12.1916 | -105.8277
12.1952 | -105.842
12.1989 | -105.8549
12.2026 | -105.8682
12.2063 | -105.8835
12.2099 | -105.8975
12.2136 | -105.9084
12.2173 | -105.9182
12.2209 | -105.9214
12.2246 | -105.9234
12.2283 | -105.9237
12.232 | -105.9205
12.2356 | -105.9154
12.2393 | -105.908
12.243 | -105.8979
12.2467 | -105.8863
12.2503 | -105.8708
12.254 | -105.8566
12.2577 | -105.8408
12.2614 | -105.8252
12.265 | -105.812
12.2687 | -105.7977
12.2724 | -105.7824
12.276 | -105.768
12.2797 | -105.7516
12.2834 | -105.7347
12.2871 | -105.7239
12.2907 | -105.7121
12.2944 | -105.6986
12.2981 | -105.6875
12.3018 | -105.6758
12.3054 | -105.6669
12.3091 | -105.6617
12.3128 | -105.656
12.3165 | -105.6486
12.3201 | -105.6441
12.3238 | -105.644
12.3275 | -105.6405
12.3311 | -105.6368
12.3348 | -105.6311
12.3385 | -105.6281
12.3422 | -105.6249
12.3458 | -105.6217
12.3495 | -105.62
12.3532 | -105.6187
12.3569 | -105.6163
12.3605 | -105.6175
12.3642 | -105.6187
12.3679 | -105.6163
12.3716 | -105.6171
12.3752 | -105.6185
12.3789 | -105.6203
12.3826 | -105.6238
12.3862 | -105.6284
12.3899 | -105.6359
12.3936 | -105.6435
12.3973 | -105.6498
12.4009 | -105.6593
12.4046 | -105.6676
12.4083 | -105.6737
12.412 | -105.6827
12.4156 | -105.6905
12.4193 | -105.6958
12.423 | -105.7032
12.4267 | -105.7095
12.4303 | -105.7139
12.434 | -105.7165
12.4377 | -105.719
12.4413 | -105.7252
12.445 | -105.7301
12.4487 | -105.7321
12.4524 | -105.7361
12.456 | -105.7387
12.4597 | -105.7405
12.4634 | -105.7436
12.4671 | -105.7426
12.4707 | -105.7466
12.4744 | -105.7547
12.4781 | -105.7599
12.4817 | -105.7637
12.4854 | -105.7682
12.4891 | -105.7682
12.4928 | -105.7717
12.4964 | -105.7779
12.5001 | -105.7851
12.5038 | -105.7917
12.5075 | -105.7981
12.5111 | -105.805
12.5148 | -105.8096
12.5185 | -105.8162
12.5222 | -105.8208
12.5258 | -105.8254
12.5295 | -105.827
12.5332 | -105.8275
12.5368 | -105.8303
12.5405 | -105.8341
12.5442 | -105.8337
12.5479 | -105.8304
12.5515 | -105.8217
12.5552 | -105.8104
12.5589 | -105.7982
12.5626 | -105.7853
12.5662 | -105.7707
12.5699 | -105.7553
12.5736 | -105.7396
12.5773 | -105.7258
12.5809 | -105.7113
12.5846 | -105.6953
12.5883 | -105.6798
12.5919 | -105.6624
12.5956 | -105.6467
12.5993 | -105.6292
12.603 | -105.6128
12.6066 | -105.5974
12.6103 | -105.588
12.614 | -105.5787
12.6177 | -105.5701
12.6213 | -105.5623
12.625 | -105.5544
12.6287 | -105.5499
12.6324 | -105.5474
12.636 | -105.5481
12.6397 | -105.5507
12.6434 | -105.5539
12.647 | -105.5604
12.6507 | -105.564
12.6544 | -105.5682
12.6581 | -105.577
12.6617 | -105.5892
12.6654 | -105.5993
12.6691 | -105.6103
12.6728 | -105.6212
12.6764 | -105.6336
12.6801 | -105.6439
12.6838 | -105.6506
12.6875 | -105.6571
12.6911 | -105.6647
12.6948 | -105.6743
12.6985 | -105.6792
12.7021 | -105.6834
12.7058 | -105.6872
12.7095 | -105.6879
12.7132 | -105.6814
12.7168 | -105.6726
12.7205 | -105.668
12.7242 | -105.6629
12.7279 | -105.661
12.7315 | -105.6568
12.7352 | -105.65
12.7389 | -105.6444
12.7425 | -105.6389
12.7462 | -105.6319
12.7499 | -105.624
12.7536 | -105.614
12.7572 | -105.6037
12.7609 | -105.5918
12.7646 | -105.5798
12.7683 | -105.5692
12.7719 | -105.5614
12.7756 | -105.555
12.7793 | -105.5488
12.783 | -105.5443
12.7866 | -105.5406
12.7903 | -105.5371
12.794 | -105.5335
12.7976 | -105.5332
12.8013 | -105.5353
12.805 | -105.5378
12.8087 | -105.5442
12.8123 | -105.5505
12.816 | -105.5554
12.8197 | -105.5621
12.8234 | -105.5686
12.827 | -105.5765
12.8307 | -105.5885
12.8344 | -105.6016
12.8381 | -105.6098
12.8417 | -105.6159
12.8454 | -105.6246
12.8491 | -105.6323
12.8527 | -105.6385
12.8564 | -105.6407
12.8601 | -105.6407
12.8638 | -105.64
12.8674 | -105.6396
12.8711 | -105.6407
12.8748 | -105.6373
12.8785 | -105.6301
12.8821 | -105.6209
12.8858 | -105.6088
12.8895 | -105.5915
12.8932 | -105.5763
12.8968 | -105.5649
12.9005 | -105.5522
12.9042 | -105.5328
12.9078 | -105.5106
12.9115 | -105.4908
12.9152 | -105.4755
12.9189 | -105.4626
12.9225 | -105.4515
12.9262 | -105.441
12.9299 | -105.4312
12.9336 | -105.4229
12.9372 | -105.4181
12.9409 | -105.4167
12.9446 | -105.4161
12.9483 | -105.4152
12.9519 | -105.4149
12.9556 | -105.4175
12.9593 | -105.4162
12.9629 | -105.4188
12.9666 | -105.4236
12.9703 | -105.4279
12.974 | -105.4297
12.9776 | -105.4337
12.9813 | -105.4358
12.985 | -105.4361
12.9887 | -105.437
12.9923 | -105.4387
12.996 | -105.4396
12.9997 | -105.4431
13.0034 | -105.4459
13.007 | -105.4454
13.0107 | -105.4454
13.0144 | -105.4446
13.018 | -105.4411
13.0217 | -105.4426
13.0254 | -105.4489
13.0291 | -105.4572
13.0327 | -105.4614
13.0364 | -105.4668
13.0401 | -105.4748
13.0438 | -105.4864
13.0474 | -105.4985
13.0511 | -105.5122
13.0548 | -105.5261
13.0585 | -105.5378
13.0621 | -105.5525
13.0658 | -105.5711
13.0695 | -105.5893
13.0731 | -105.607
13.0768 | -105.628
13.0805 | -105.6459
13.0842 | -105.6591
13.0878 | -105.6697
13.0915 | -105.6856
13.0952 | -105.7045
13.0989 | -105.7167
13.1025 | -105.7254
13.1062 | -105.7361
13.1099 | -105.7479
13.1136 | -105.7587
13.1172 | -105.7696
13.1209 | -105.7804
13.1246 | -105.7884
13.1282 | -105.7967
13.1319 | -105.8066
13.1356 | -105.8164
13.1393 | -105.8274
13.1429 | -105.836
13.1466 | -105.8455
13.1503 | -105.8521
13.154 | -105.8626
13.1576 | -105.874
13.1613 | -105.8854
13.165 | -105.8959
13.1686 | -105.9071
13.1723 | -105.9181
13.176 | -105.9275
13.1797 | -105.9361
13.1833 | -105.9423
13.187 | -105.9471
13.1907 | -105.9513
13.1944 | -105.9541
13.198 | -105.9553
13.2017 | -105.9542
13.2054 | -105.949
13.2091 | -105.9389
13.2127 | -105.9265
13.2164 | -105.9096
13.2201 | -105.8917
13.2237 | -105.8732
13.2274 | -105.8529
13.2311 | -105.8265
13.2348 | -105.7967
13.2384 | -105.7691
13.2421 | -105.7453
13.2458 | -105.7245
13.2495 | -105.7017
13.2531 | -105.6798
13.2568 | -105.6585
13.2605 | -105.6389
13.2642 | -105.6241
13.2678 | -105.6117
13.2715 | -105.5995
13.2752 | -105.5911
13.2788 | -105.5849
13.2825 | -105.5793
13.2862 | -105.5773
13.2899 | -105.58
13.2935 | -105.5825
13.2972 | -105.5852
13.3009 | -105.591
13.3046 | -105.5977
13.3082 | -105.607
13.3119 | -105.6181
13.3156 | -105.6322
13.3193 | -105.6466
13.3229 | -105.6562
13.3266 | -105.6675
13.3303 | -105.6796
13.3339 | -105.6947
13.3376 | -105.7101
13.3413 | -105.7256
13.345 | -105.7358
13.3486 | -105.7474
13.3523 | -105.7579
13.356 | -105.7692
13.3597 | -105.7817
13.3633 | -105.793
13.367 | -105.8025
13.3707 | -105.8092
13.3744 | -105.817
13.378 | -105.8204
13.3817 | -105.8249
13.3854 | -105.8294
13.389 | -105.8321
13.3927 | -105.8327
13.3964 | -105.835
13.4001 | -105.8351
13.4037 | -105.8328
13.4074 | -105.8308
13.4111 | -105.8337
13.4148 | -105.8357
13.4184 | -105.8356
13.4221 | -105.8376
13.4258 | -105.8423
13.4294 | -105.8472
13.4331 | -105.847
13.4368 | -105.8442
13.4405 | -105.8451
13.4441 | -105.8509
13.4478 | -105.8519
13.4515 | -105.8532
13.4552 | -105.8568
13.4588 | -105.8587
13.4625 | -105.8529
13.4662 | -105.8484
13.4699 | -105.8455
13.4735 | -105.8399
13.4772 | -105.8346
13.4809 | -105.8256
13.4845 | -105.8128
13.4882 | -105.8024
13.4919 | -105.7928
13.4956 | -105.7825
13.4992 | -105.7735
13.5029 | -105.7631
13.5066 | -105.7529
13.5103 | -105.7421
13.5139 | -105.7326
13.5176 | -105.7283
13.5213 | -105.7217
13.525 | -105.7142
13.5286 | -105.7085
13.5323 | -105.7058
13.536 | -105.7038
13.5396 | -105.7048
13.5433 | -105.7103
13.547 | -105.717
13.5507 | -105.7213
13.5543 | -105.7248
13.558 | -105.7315
13.5617 | -105.7419
13.5654 | -105.7551
13.569 | -105.771
13.5727 | -105.7848
13.5764 | -105.798
13.5801 | -105.8122
13.5837 | -105.8245
13.5874 | -105.8382
13.5911 | -105.8502
13.5947 | -105.8607
13.5984 | -105.872
13.6021 | -105.8811
13.6058 | -105.8823
13.6094 | -105.8854
13.6131 | -105.8879
13.6168 | -105.8867
13.6205 | -105.8833
13.6241 | -105.8737
13.6278 | -105.8683
13.6315 | -105.8668
13.6352 | -105.8577
13.6388 | -105.8449
13.6425 | -105.8355
13.6462 | -105.8237
13.6498 | -105.8131
13.6535 | -105.8052
13.6572 | -105.8017
13.6609 | -105.8001
13.6645 | -105.7995
13.6682 | -105.8004
13.6719 | -105.8
13.6756 | -105.8022
13.6792 | -105.808
13.6829 | -105.8159
13.6866 | -105.8251
13.6903 | -105.8338
13.6939 | -105.8448
13.6976 | -105.8563
13.7013 | -105.8677
13.7049 | -105.8819
13.7086 | -105.8964
13.7123 | -105.9127
13.716 | -105.9303
13.7196 | -105.946
13.7233 | -105.9576
13.727 | -105.9698
13.7307 | -105.9842
13.7343 | -105.9975
13.738 | -106.0091
13.7417 | -106.0224
13.7454 | -106.0335
13.749 | -106.0468
13.7527 | -106.0559
13.7564 | -106.0618
13.76 | -106.068
13.7637 | -106.0715
13.7674 | -106.0787
13.7711 | -106.0815
13.7747 | -106.0776
13.7784 | -106.0784
13.7821 | -106.08
13.7858 | -106.0804
13.7894 | -106.0809
13.7931 | -106.0787
13.7968 | -106.0734
13.8004 | -106.0685
13.8041 | -106.0636
13.8078 | -106.0521
13.8115 | -106.0391
13.8151 | -106.0294
13.8188 | -106.0176
13.8225 | -106.007
13.8262 | -105.9928
13.8298 | -105.9755
13.8335 | -105.9614
13.8372 | -105.9491
13.8409 | -105.9362
13.8445 | -105.9228
13.8482 | -105.9094
13.8519 | -105.893
13.8555 | -105.8775
13.8592 | -105.8601
13.8629 | -105.8427
13.8666 | -105.8285
13.8702 | -105.8167
13.8739 | -105.8107
13.8776 | -105.8073
13.8813 | -105.801
13.8849 | -105.7964
13.8886 | -105.7954
13.8923 | -105.7938
13.896 | -105.7949
13.8996 | -105.7984
13.9033 | -105.8024
13.907 | -105.8067
13.9106 | -105.811
13.9143 | -105.8145
13.918 | -105.8211
13.9217 | -105.8297
13.9253 | -105.8349
13.929 | -105.8408
13.9327 | -105.8505
13.9364 | -105.8552
13.94 | -105.8593
13.9437 | -105.8662
13.9474 | -105.869
13.9511 | -105.8763
13.9547 | -105.8815
13.9584 | -105.882
13.9621 | -105.8814
13.9657 | -105.8816
13.9694 | -105.8863
13.9731 | -105.8894
13.9768 | -105.8876
13.9804 | -105.8822
13.9841 | -105.8782
13.9878 | -105.8751
13.9915 | -105.8721
13.9951 | -105.8701
13.9988 | -105.8672
14.0025 | -105.8614
14.0062 | -105.8551
14.0098 | -105.8506
14.0135 | -105.8446
14.0172 | -105.8317
14.0208 | -105.8234
14.0245 | -105.8175
14.0282 | -105.8113
14.0319 | -105.8053
14.0355 | -105.8002
14.0392 | -105.7957
14.0429 | -105.7952
14.0466 | -105.796
14.0502 | -105.7979
14.0539 | -105.8041
14.0576 | -105.8132
14.0613 | -105.8242
14.0649 | -105.8348
14.0686 | -105.8462
14.0723 | -105.8591
14.0759 | -105.8693
14.0796 | -105.8799
14.0833 | -105.8933
14.087 | -105.9045
14.0906 | -105.9137
14.0943 | -105.9285
14.098 | -105.9394
14.1017 | -105.9482
14.1053 | -105.9558
14.109 | -105.9645
14.1127 | -105.9723
14.1164 | -105.9808
14.12 | -105.9851
14.1237 | -105.9885
14.1274 | -105.9948
14.131 | -105.998
14.1347 | -105.9958
14.1384 | -105.9964
14.1421 | -105.9936
14.1457 | -105.9907
14.1494 | -105.9896
14.1531 | -105.9846
14.1568 | -105.9808
14.1604 | -105.9784
14.1641 | -105.9759
14.1678 | -105.9721
14.1715 | -105.9668
14.1751 | -105.958
14.1788 | -105.9487
14.1825 | -105.9428
14.1861 | -105.9341
14.1898 | -105.9207
14.1935 | -105.9078
14.1972 | -105.8939
14.2008 | -105.8834
14.2045 | -105.8738
14.2082 | -105.8618
14.2119 | -105.8492
14.2155 | -105.8388
14.2192 | -105.8325
14.2229 | -105.821
14.2265 | -105.8081
14.2302 | -105.7935
14.2339 | -105.7799
14.2376 | -105.7676
14.2412 | -105.7557
14.2449 | -105.7418
14.2486 | -105.7315
14.2523 | -105.7198
14.2559 | -105.7073
14.2596 | -105.6968
14.2633 | -105.6879
14.267 | -105.6829
14.2706 | -105.6786
14.2743 | -105.6741
14.278 | -105.6731
14.2816 | -105.6742
14.2853 | -105.6778
14.289 | -105.6869
14.2927 | -105.698
14.2963 | -105.7072
14.3 | -105.7197
14.3037 | -105.7358
14.3074 | -105.7587
14.311 | -105.7765
14.3147 | -105.7939
14.3184 | -105.8155
14.3221 | -105.8355
14.3257 | -105.8583
14.3294 | -105.8807
14.3331 | -105.9045
14.3367 | -105.9293
14.3404 | -105.9541
14.3441 | -105.9782
14.3478 | -106.0027
14.3514 | -106.0253
14.3551 | -106.0462
14.3588 | -106.0658
14.3625 | -106.0852
14.3661 | -106.1058
14.3698 | -106.1247
14.3735 | -106.1411
14.3772 | -106.1555
14.3808 | -106.1728
14.3845 | -106.1925
14.3882 | -106.2069
14.3918 | -106.2198
14.3955 | -106.2355
14.3992 | -106.2486
14.4029 | -106.2613
14.4065 | -106.2691
14.4102 | -106.2752
14.4139 | -106.2818
14.4176 | -106.2852
14.4212 | -106.2884
14.4249 | -106.2878
14.4286 | -106.2847
14.4323 | -106.2811
14.4359 | -106.2796
14.4396 | -106.2776
14.4433 | -106.2754
14.4469 | -106.2718
14.4506 | -106.266
14.4543 | -106.2588
14.458 | -106.2536
14.4616 | -106.2468
14.4653 | -106.2387
14.469 | -106.2314
14.4727 | -106.2239
14.4763 | -106.2181
14.48 | -106.2122
14.4837 | -106.2036
14.4873 | -106.1932
14.491 | -106.1823
14.4947 | -106.1749
14.4984 | -106.1693
14.502 | -106.1627
14.5057 | -106.158
14.5094 | -106.1516
14.5131 | -106.1459
14.5167 | -106.1407
14.5204 | -106.1341
14.5241 | -106.1305
14.5278 | -106.1248
14.5314 | -106.1224
14.5351 | -106.1213
14.5388 | -106.1242
14.5424 | -106.1237
14.5461 | -106.1247
14.5498 | -106.1284
14.5535 | -106.1312
14.5571 | -106.1351
14.5608 | -106.1388
14.5645 | -106.1426
14.5682 | -106.1465
14.5718 | -106.1486
14.5755 | -106.1503
14.5792 | -106.1524
14.5829 | -106.1536
14.5865 | -106.1535
14.5902 | -106.1521
14.5939 | -106.1506
14.5975 | -106.1529
14.6012 | -106.153
14.6049 | -106.1489
14.6086 | -106.1448
14.6122 | -106.1421
14.6159 | -106.1377
14.6196 | -106.1335
14.6233 | -106.131
14.6269 | -106.129
14.6306 | -106.1252
14.6343 | -106.1219
14.638 | -106.1206
14.6416 | -106.1194
14.6453 | -106.1178
14.649 | -106.1149
14.6526 | -106.1114
14.6563 | -106.1094
14.66 | -106.1109
14.6637 | -106.113
14.6673 | -106.1153
14.671 | -106.12
14.6747 | -106.1233
14.6784 | -106.1256
14.682 | -106.1294
14.6857 | -106.1343
14.6894 | -106.1415
14.6931 | -106.1495
14.6967 | -106.159
14.7004 | -106.1681
14.7041 | -106.1715
14.7077 | -106.1742
14.7114 | -106.1777
14.7151 | -106.1826
14.7188 | -106.1895
14.7224 | -106.1956
14.7261 | -106.1981
14.7298 | -106.1996
14.7335 | -106.1994
14.7371 | -106.2014
14.7408 | -106.2024
14.7445 | -106.2008
14.7482 | -106.198
14.7518 | -106.1978
14.7555 | -106.1973
14.7592 | -106.1969
14.7628 | -106.2
14.7665 | -106.2003
14.7702 | -106.1975
14.7739 | -106.1962
14.7775 | -106.1963
14.7812 | -106.1987
14.7849 | -106.1966
14.7886 | -106.1954
14.7922 | -106.1969
14.7959 | -106.1983
14.7996 | -106.2013
14.8033 | -106.2016
14.8069 | -106.1986
14.8106 | -106.1973
14.8143 | -106.1928
14.8179 | -106.1898
14.8216 | -106.1862
14.8253 | -106.1847
14.829 | -106.1825
14.8326 | -106.1802
14.8363 | -106.1758
14.84 | -106.1717
14.8437 | -106.1652
14.8473 | -106.1592
14.851 | -106.1549
14.8547 | -106.1527
14.8583 | -106.1483
14.862 | -106.1425
14.8657 | -106.138
14.8694 | -106.1342
14.873 | -106.13
14.8767 | -106.1221
14.8804 | -106.1138
14.8841 | -106.1049
14.8877 | -106.0994
14.8914 | -106.0962
14.8951 | -106.0925
14.8988 | -106.0881
14.9024 | -106.0851
14.9061 | -106.0811
14.9098 | -106.0783
14.9134 | -106.0767
14.9171 | -106.0746
14.9208 | -106.0711
14.9245 | -106.0706
14.9281 | -106.0687
14.9318 | -106.0661
14.9355 | -106.0656
14.9392 | -106.0652
14.9428 | -106.0648
14.9465 | -106.0648
14.9502 | -106.064
14.9539 | -106.0653
14.9575 | -106.0673
14.9612 | -106.0689
14.9649 | -106.0711
14.9685 | -106.0711
14.9722 | -106.0692
14.9759 | -106.0687
14.9796 | -106.0701
14.9832 | -106.0705
14.9869 | -106.0701
14.9906 | -106.07
14.9943 | -106.0678
14.9979 | -106.0697
15.0016 | -106.07
15.0053 | -106.0701
15.009 | -106.0709
15.0126 | -106.07
15.0163 | -106.069
15.02 | -106.0672
15.0236 | -106.0657
15.0273 | -106.0704
15.031 | -106.0752
15.0347 | -106.0797
15.0383 | -106.083
15.042 | -106.0869
15.0457 | -106.0906
15.0494 | -106.0911
15.053 | -106.096
15.0567 | -106.1023
15.0604 | -106.1112
15.0641 | -106.1173
15.0677 | -106.1219
15.0714 | -106.1278
15.0751 | -106.1353
15.0787 | -106.1458
15.0824 | -106.1563
15.0861 | -106.1651
15.0898 | -106.1739
15.0934 | -106.1841
15.0971 | -106.1942
15.1008 | -106.2026
15.1045 | -106.2077
15.1081 | -106.2136
15.1118 | -106.2194
15.1155 | -106.224
15.1192 | -106.2309
15.1228 | -106.2404
15.1265 | -106.2484
15.1302 | -106.2548
15.1338 | -106.262
15.1375 | -106.2685
15.1412 | -106.2713
15.1449 | -106.2766
15.1485 | -106.2803
15.1522 | -106.2817
15.1559 | -106.2829
15.1596 | -106.2833
15.1632 | -106.2821
15.1669 | -106.2791
15.1706 | -106.275
15.1743 | -106.2712
15.1779 | -106.2688
15.1816 | -106.2646
15.1853 | -106.2615
15.1889 | -106.262
15.1926 | -106.2577
15.1963 | -106.2522
15.2 | -106.2487
15.2036 | -106.2413
15.2073 | -106.2349
15.211 | -106.2279
15.2147 | -106.2195
15.2183 | -106.2103
15.222 | -106.201
15.2257 | -106.1931
15.2293 | -106.186
15.233 | -106.181
15.2367 | -106.1776
15.2404 | -106.175
15.244 | -106.1691
15.2477 | -106.1607
15.2514 | -106.1518
15.2551 | -106.1458
15.2587 | -106.1392
15.2624 | -106.1323
15.2661 | -106.1295
15.2698 | -106.1238
15.2734 | -106.1166
15.2771 | -106.1076
15.2808 | -106.102
15.2844 | -106.0969
15.2881 | -106.0883
15.2918 | -106.0798
15.2955 | -106.0696
15.2991 | -106.0604
15.3028 | -106.052
15.3065 | -106.0441
15.3102 | -106.0362
15.3138 | -106.0248
15.3175 | -106.0162
15.3212 | -106.009
15.3249 | -106.001
15.3285 | -105.9922
15.3322 | -105.9858
15.3359 | -105.9793
15.3395 | -105.972
15.3432 | -105.9656
15.3469 | -105.9615
15.3506 | -105.9596
15.3542 | -105.9599
15.3579 | -105.9596
15.3616 | -105.9589
15.3653 | -105.9652
15.3689 | -105.9657
15.3726 | -105.969
15.3763 | -105.9741
15.38 | -105.9811
15.3836 | -105.9892
15.3873 | -105.9948
15.391 | -106.0034
15.3946 | -106.0146
15.3983 | -106.0262
15.402 | -106.0399
15.4057 | -106.0571
15.4093 | -106.073
15.413 | -106.0906
15.4167 | -106.1094
15.4204 | -106.1299
15.424 | -106.1519
15.4277 | -106.1739
15.4314 | -106.1974
15.4351 | -106.2212
15.4387 | -106.2457
15.4424 | -106.266
15.4461 | -106.2844
15.4497 | -106.3046
15.4534 | -106.3206
15.4571 | -106.3368
15.4608 | -106.354
15.4644 | -106.3725
15.4681 | -106.3896
15.4718 | -106.4055
15.4755 | -106.4194
15.4791 | -106.436
15.4828 | -106.4492
15.4865 | -106.46
15.4902 | -106.47
15.4938 | -106.477
15.4975 | -106.482
15.5012 | -106.4879
15.5048 | -106.4916
15.5085 | -106.4956
15.5122 | -106.5005
15.5159 | -106.5045
15.5195 | -106.5023
15.5232 | -106.5005
15.5269 | -106.4978
15.5306 | -106.495
15.5342 | -106.4927
15.5379 | -106.4867
15.5416 | -106.4789
15.5452 | -106.4692
15.5489 | -106.4624
15.5526 | -106.4541
15.5563 | -106.4439
15.5599 | -106.4357
15.5636 | -106.4275
15.5673 | -106.4152
15.571 | -106.4024
15.5746 | -106.3913
15.5783 | -106.3812
15.582 | -106.3692
15.5857 | -106.356
15.5893 | -106.3385
15.593 | -106.3211
15.5967 | -106.2997
15.6003 | -106.2817
15.604 | -106.2626
15.6077 | -106.2412
15.6114 | -106.2228
15.615 | -106.2037
15.6187 | -106.184
15.6224 | -106.1625
15.6261 | -106.139
15.6297 | -106.116
15.6334 | -106.0961
15.6371 | -106.0769
15.6408 | -106.0582
15.6444 | -106.039
15.6481 | -106.0222
15.6518 | -106.0075
15.6554 | -105.9934
15.6591 | -105.9782
15.6628 | -105.9661
15.6665 | -105.9573
15.6701 | -105.9498
15.6738 | -105.9419
15.6775 | -105.9381
15.6812 | -105.9343
15.6848 | -105.9316
15.6885 | -105.9278
15.6922 | -105.9253
15.6959 | -105.9242
15.6995 | -105.924
15.7032 | -105.924
15.7069 | -105.9255
15.7105 | -105.926
15.7142 | -105.9263
15.7179 | -105.9318
15.7216 | -105.9365
15.7252 | -105.9381
15.7289 | -105.9419
15.7326 | -105.9456
15.7363 | -105.9493
15.7399 | -105.9525
15.7436 | -105.9579
15.7473 | -105.9622
15.751 | -105.9655
15.7546 | -105.9673
15.7583 | -105.9643
15.762 | -105.964
15.7656 | -105.9624
15.7693 | -105.9606
15.773 | -105.9589
15.7767 | -105.9579
15.7803 | -105.954
15.784 | -105.9476
15.7877 | -105.9406
15.7914 | -105.9363
15.795 | -105.9319
15.7987 | -105.9275
15.8024 | -105.9225
15.8061 | -105.9194
15.8097 | -105.9175
15.8134 | -105.9164
15.8171 | -105.9134
15.8207 | -105.9139
15.8244 | -105.9099
15.8281 | -105.9049
15.8318 | -105.8983
15.8354 | -105.8936
15.8391 | -105.8905
15.8428 | -105.8871
15.8465 | -105.8861
15.8501 | -105.888
15.8538 | -105.8858
15.8575 | -105.8823
15.8612 | -105.8822
15.8648 | -105.8844
15.8685 | -105.8858
15.8722 | -105.8878
15.8758 | -105.8877
15.8795 | -105.8886
15.8832 | -105.8926
15.8869 | -105.894
15.8905 | -105.8968
15.8942 | -105.8995
15.8979 | -105.9021
15.9016 | -105.9013
15.9052 | -105.8999
15.9089 | -105.8971
15.9126 | -105.8945
15.9162 | -105.8952
15.9199 | -105.8959
15.9236 | -105.8966
15.9273 | -105.8938
15.9309 | -105.8912
15.9346 | -105.8887
15.9383 | -105.8863
15.942 | -105.8841
15.9456 | -105.8778
15.9493 | -105.8728
15.953 | -105.8676
15.9567 | -105.8582
15.9603 | -105.8479
15.964 | -105.838
15.9677 | -105.8299
15.9713 | -105.822
15.975 | -105.815
15.9787 | -105.8059
15.9824 | -105.7978
15.986 | -105.7919
15.9897 | -105.7823
15.9934 | -105.7747
15.9971 | -105.7678
16.0007 | -105.7598
16.0044 | -105.753
16.0081 | -105.7482
16.0118 | -105.7464
16.0154 | -105.7444
16.0191 | -105.7423
16.0228 | -105.7409
16.0264 | -105.74
16.0301 | -105.7424
16.0338 | -105.7451
16.0375 | -105.7499
16.0411 | -105.7574
16.0448 | -105.7658
16.0485 | -105.7719
16.0522 | -105.781
16.0558 | -105.7933
16.0595 | -105.8081
16.0632 | -105.8204
16.0669 | -105.8296
16.0705 | -105.841
16.0742 | -105.853
16.0779 | -105.8651
16.0815 | -105.8802
16.0852 | -105.8963
16.0889 | -105.9104
16.0926 | -105.9227
16.0962 | -105.9367
16.0999 | -105.9499
16.1036 | -105.9641
16.1073 | -105.9797
16.1109 | -105.9943
16.1146 | -106.0076
16.1183 | -106.0182
16.122 | -106.0307
16.1256 | -106.0432
16.1293 | -106.0567
16.133 | -106.0706
16.1366 | -106.0837
16.1403 | -106.0972
16.144 | -106.1088
16.1477 | -106.1178
16.1513 | -106.1266
16.155 | -106.1332
16.1587 | -106.1427
16.1624 | -106.15
16.166 | -106.1587
16.1697 | -106.168
16.1734 | -106.174
16.1771 | -106.1794
16.1807 | -106.1842
16.1844 | -106.1902
16.1881 | -106.198
16.1917 | -106.2064
16.1954 | -106.2112
16.1991 | -106.2137
16.2028 | -106.2173
16.2064 | -106.2201
16.2101 | -106.2251
16.2138 | -106.2291
16.2175 | -106.2343
16.2211 | -106.2369
16.2248 | -106.2358
16.2285 | -106.2353
16.2322 | -106.2348
16.2358 | -106.2331
16.2395 | -106.2325
16.2432 | -106.2304
16.2468 | -106.2275
16.2505 | -106.2232
16.2542 | -106.2167
16.2579 | -106.2089
16.2615 | -106.2018
16.2652 | -106.195
16.2689 | -106.1895
16.2726 | -106.1844
16.2762 | -106.1796
16.2799 | -106.1753
16.2836 | -106.1696
16.2872 | -106.1626
16.2909 | -106.1553
16.2946 | -106.1522
16.2983 | -106.1496
16.3019 | -106.1468
16.3056 | -106.1452
16.3093 | -106.1441
16.313 | -106.1424
16.3166 | -106.1437
16.3203 | -106.1452
16.324 | -106.1464
16.3277 | -106.1505
16.3313 | -106.1568
16.335 | -106.1622
16.3387 | -106.1676
16.3423 | -106.1752
16.346 | -106.1855
16.3497 | -106.1952
16.3534 | -106.2009
16.357 | -106.2085
16.3607 | -106.217
16.3644 | -106.225
16.3681 | -106.2302
16.3717 | -106.2348
16.3754 | -106.2395
16.3791 | -106.2448
16.3828 | -106.2505
16.3864 | -106.2561
16.3901 | -106.258
16.3938 | -106.2585
16.3974 | -106.2595
16.4011 | -106.2585
16.4048 | -106.2568
16.4085 | -106.2563
16.4121 | -106.2552
16.4158 | -106.2519
16.4195 | -106.2476
16.4232 | -106.2423
16.4268 | -106.2413
16.4305 | -106.2412
16.4342 | -106.2417
16.4379 | -106.2401
16.4415 | -106.2382
16.4452 | -106.2358
16.4489 | -106.2348
16.4525 | -106.2368
16.4562 | -106.2419
16.4599 | -106.2494
16.4636 | -106.2554
16.4672 | -106.2623
16.4709 | -106.2682
16.4746 | -106.276
16.4783 | -106.285
16.4819 | -106.2961
16.4856 | -106.3043
16.4893 | -106.3161
16.493 | -106.3299
16.4966 | -106.3421
16.5003 | -106.3567
16.504 | -106.3701
16.5076 | -106.3851
16.5113 | -106.4014
16.515 | -106.4173
16.5187 | -106.4337
16.5223 | -106.452
16.526 | -106.4711
16.5297 | -106.4878
16.5334 | -106.5047
16.537 | -106.5186
16.5407 | -106.5322
16.5444 | -106.5444
16.5481 | -106.5554
16.5517 | -106.5645
16.5554 | -106.5736
16.5591 | -106.5819
16.5627 | -106.5868
16.5664 | -106.5893
16.5701 | -106.5942
16.5738 | -106.5972
16.5774 | -106.5992
16.5811 | -106.6
16.5848 | -106.6025
16.5885 | -106.6074
16.5921 | -106.6088
16.5958 | -106.607
16.5995 | -106.608
16.6031 | -106.6097
16.6068 | -106.6074
16.6105 | -106.6056
16.6142 | -106.6036
16.6178 | -106.6004
16.6215 | -106.5979
16.6252 | -106.596
16.6289 | -106.5968
16.6325 | -106.5956
16.6362 | -106.5957
16.6399 | -106.5955
16.6436 | -106.5959
16.6472 | -106.5966
16.6509 | -106.596
16.6546 | -106.5974
16.6582 | -106.6006
16.6619 | -106.6018
16.6656 | -106.6027
16.6693 | -106.6039
16.6729 | -106.6057
16.6766 | -106.6078
16.6803 | -106.6102
16.684 | -106.6128
16.6876 | -106.6152
16.6913 | -106.614
16.695 | -106.6132
16.6987 | -106.6136
16.7023 | -106.6171
16.706 | -106.6203
16.7097 | -106.6239
16.7133 | -106.628
16.717 | -106.6308
16.7207 | -106.6315
16.7244 | -106.634
16.728 | -106.6338
16.7317 | -106.6337
16.7354 | -106.6325
16.7391 | -106.6295
16.7427 | -106.6262
16.7464 | -106.6204
16.7501 | -106.6149
16.7538 | -106.6098
16.7574 | -106.6043
16.7611 | -106.5987
16.7648 | -106.5909
16.7684 | -106.5821
16.7721 | -106.5751
16.7758 | -106.5691
16.7795 | -106.5604
16.7831 | -106.5524
16.7868 | -106.5454
16.7905 | -106.5352
16.7942 | -106.5253
16.7978 | -106.5177
16.8015 | -106.5089
16.8052 | -106.4999
16.8089 | -106.4913
16.8125 | -106.4822
16.8162 | -106.4736
16.8199 | -106.4651
16.8235 | -106.4559
16.8272 | -106.4467
16.8309 | -106.4401
16.8346 | -106.4346
16.8382 | -106.4293
16.8419 | -106.425
16.8456 | -106.423
16.8493 | -106.4223
16.8529 | -106.4205
16.8566 | -106.4193
16.8603 | -106.42
16.864 | -106.4206
16.8676 | -106.4211
16.8713 | -106.4222
16.875 | -106.4235
16.8786 | -106.4289
16.8823 | -106.434
16.886 | -106.4394
16.8897 | -106.4474
16.8933 | -106.4551
16.897 | -106.4626
16.9007 | -106.4713
16.9044 | -106.4796
16.908 | -106.4901
16.9117 | -106.5027
16.9154 | -106.5158
16.919 | -106.5282
16.9227 | -106.5415
16.9264 | -106.5549
16.9301 | -106.5681
16.9337 | -106.5811
16.9374 | -106.5933
16.9411 | -106.6032
16.9448 | -106.6145
16.9484 | -106.6294
16.9521 | -106.6407
16.9558 | -106.6512
16.9595 | -106.6611
16.9631 | -106.6714
16.9668 | -106.6824
16.9705 | -106.6936
16.9741 | -106.701
16.9778 | -106.7083
16.9815 | -106.716
16.9852 | -106.7246
16.9888 | -106.7325
16.9925 | -106.7362
16.9962 | -106.742
16.9999 | -106.7451
17.0035 | -106.7465
17.0072 | -106.7489
17.0109 | -106.7481
17.0146 | -106.7438
17.0182 | -106.7408
17.0219 | -106.7388
17.0256 | -106.7351
17.0292 | -106.7303
17.0329 | -106.7239
17.0366 | -106.7169
17.0403 | -106.7093
17.0439 | -106.7
17.0476 | -106.6898
17.0513 | -106.6772
17.055 | -106.6646
17.0586 | -106.6489
17.0623 | -106.6352
17.066 | -106.6218
17.0697 | -106.6091
17.0733 | -106.5956
17.077 | -106.5824
17.0807 | -106.5673
17.0843 | -106.5509
17.088 | -106.5361
17.0917 | -106.5179
17.0954 | -106.5038
17.099 | -106.4923
17.1027 | -106.4802
17.1064 | -106.4708
17.1101 | -106.461
17.1137 | -106.4528
17.1174 | -106.4467
17.1211 | -106.4396
17.1248 | -106.434
17.1284 | -106.4303
17.1321 | -106.4274
17.1358 | -106.4248
17.1394 | -106.4236
17.1431 | -106.4255
17.1468 | -106.43
17.1505 | -106.4356
17.1541 | -106.4403
17.1578 | -106.4464
17.1615 | -106.4533
17.1652 | -106.4621
17.1688 | -106.4731
17.1725 | -106.4868
17.1762 | -106.5015
17.1799 | -106.5158
17.1835 | -106.5307
17.1872 | -106.5461
17.1909 | -106.5598
17.1945 | -106.5748
17.1982 | -106.5902
17.2019 | -106.6056
17.2056 | -106.622
17.2092 | -106.6399
17.2129 | -106.6559
17.2166 | -106.6673
17.2203 | -106.6787
17.2239 | -106.6885
17.2276 | -106.6972
17.2313 | -106.7039
17.235 | -106.7102
17.2386 | -106.7169
17.2423 | -106.7196
17.246 | -106.7229
17.2496 | -106.723
17.2533 | -106.7234
17.257 | -106.723
17.2607 | -106.7215
17.2643 | -106.7159
17.268 | -106.7095
17.2717 | -106.7031
17.2754 | -106.6944
17.279 | -106.6854
17.2827 | -106.6757
17.2864 | -106.6656
17.2901 | -106.6548
17.2937 | -106.6427
17.2974 | -106.6297
17.3011 | -106.6189
17.3047 | -106.6099
17.3084 | -106.5976
17.3121 | -106.5852
17.3158 | -106.5731
17.3194 | -106.5624
17.3231 | -106.554
17.3268 | -106.5476
17.3305 | -106.5437
17.3341 | -106.5405
17.3378 | -106.5388
17.3415 | -106.5367
17.3451 | -106.534
17.3488 | -106.5342
17.3525 | -106.5364
17.3562 | -106.5413
17.3598 | -106.5457
17.3635 | -106.5529
17.3672 | -106.563
17.3709 | -106.576
17.3745 | -106.5901
17.3782 | -106.6029
17.3819 | -106.6194
17.3856 | -106.6377
17.3892 | -106.6557
17.3929 | -106.6763
17.3966 | -106.6968
17.4002 | -106.7173
17.4039 | -106.736
17.4076 | -106.7544
17.4113 | -106.7713
17.4149 | -106.7901
17.4186 | -106.8072
17.4223 | -106.8241
17.426 | -106.8418
17.4296 | -106.8576
17.4333 | -106.8717
17.437 | -106.888
17.4407 | -106.9023
17.4443 | -106.9134
17.448 | -106.9231
17.4517 | -106.9321
17.4553 | -106.9387
17.459 | -106.9444
17.4627 | -106.9489
17.4664 | -106.9542
17.47 | -106.9612
17.4737 | -106.9651
17.4774 | -106.9658
17.4811 | -106.9641
17.4847 | -106.9559
17.4884 | -106.9528
17.4921 | -106.9476
17.4958 | -106.9425
17.4994 | -106.9382
17.5031 | -106.9307
17.5068 | -106.9235
17.5104 | -106.9157
17.5141 | -106.9072
17.5178 | -106.8984
17.5215 | -106.8885
17.5251 | -106.8793
17.5288 | -106.8715
17.5325 | -106.8648
17.5362 | -106.8569
17.5398 | -106.8496
17.5435 | -106.8437
17.5472 | -106.8379
17.5509 | -106.8304
17.5545 | -106.8228
17.5582 | -106.8204
17.5619 | -106.8177
17.5655 | -106.8122
17.5692 | -106.8079
17.5729 | -106.8045
17.5766 | -106.8016
17.5802 | -106.7967
17.5839 | -106.7911
17.5876 | -106.7877
17.5913 | -106.7853
17.5949 | -106.7836
17.5986 | -106.7816
17.6023 | -106.778
17.606 | -106.7732
17.6096 | -106.7704
17.6133 | -106.7658
17.617 | -106.76
17.6206 | -106.7523
17.6243 | -106.7456
17.628 | -106.7428
17.6317 | -106.7371
17.6353 | -106.7295
17.639 | -106.7216
17.6427 | -106.7133
17.6464 | -106.7051
17.65 | -106.6965
17.6537 | -106.6907
17.6574 | -106.685
17.661 | -106.6768
17.6647 | -106.6667
17.6684 | -106.6573
17.6721 | -106.6462
17.6757 | -106.6352
17.6794 | -106.6268
17.6831 | -106.616
17.6868 | -106.6042
17.6904 | -106.5947
17.6941 | -106.585
17.6978 | -106.5756
17.7015 | -106.5664
17.7051 | -106.5593
17.7088 | -106.5523
17.7125 | -106.5469
17.7161 | -106.5407
17.7198 | -106.5333
17.7235 | -106.5283
17.7272 | -106.5271
17.7308 | -106.5265
17.7345 | -106.5241
17.7382 | -106.5219
17.7419 | -106.5219
17.7455 | -106.5213
17.7492 | -106.5205
17.7529 | -106.5193
17.7566 | -106.5212
17.7602 | -106.5249
17.7639 | -106.5267
17.7676 | -106.5273
17.7712 | -106.5263
17.7749 | -106.5255
17.7786 | -106.5252
17.7823 | -106.5258
17.7859 | -106.526
17.7896 | -106.5275
17.7933 | -106.5294
17.797 | -106.53
17.8006 | -106.532
17.8043 | -106.5322
17.808 | -106.5341
17.8117 | -106.5329
17.8153 | -106.5316
17.819 | -106.5337
17.8227 | -106.5351
17.8263 | -106.5351
17.83 | -106.5324
17.8337 | -106.5336
17.8374 | -106.5328
17.841 | -106.5295
17.8447 | -106.5247
17.8484 | -106.5197
17.8521 | -106.5151
17.8557 | -106.5105
17.8594 | -106.5049
17.8631 | -106.4993
17.8668 | -106.4931
17.8704 | -106.485
17.8741 | -106.4755
17.8778 | -106.4673
17.8814 | -106.4595
17.8851 | -106.4519
17.8888 | -106.4455
17.8925 | -106.4383
17.8961 | -106.4299
17.8998 | -106.4227
17.9035 | -106.4155
17.9072 | -106.408
17.9108 | -106.4013
17.9145 | -106.3955
17.9182 | -106.3913
17.9219 | -106.3878
17.9255 | -106.3859
17.9292 | -106.3845
17.9329 | -106.3822
17.9365 | -106.3841
17.9402 | -106.3849
17.9439 | -106.3833
17.9476 | -106.386
17.9512 | -106.3871
17.9549 | -106.3878
17.9586 | -106.3893
17.9623 | -106.3889
17.9659 | -106.3914
17.9696 | -106.3935
17.9733 | -106.392
17.9769 | -106.3917
17.9806 | -106.392
17.9843 | -106.3925
17.988 | -106.3941
17.9916 | -106.3946
17.9953 | -106.3934
17.999 | -106.3928
18.0027 | -106.3899
18.0063 | -106.3869
18.01 | -106.3797
18.0137 | -106.3735
18.0174 | -106.3639
18.021 | -106.3581
18.0247 | -106.3529
18.0284 | -106.3473
18.032 | -106.3414
18.0357 | -106.3317
18.0394 | -106.3198
18.0431 | -106.3081
18.0467 | -106.297
18.0504 | -106.2857
18.0541 | -106.2736
18.0578 | -106.2615
18.0614 | -106.2533
18.0651 | -106.2429
18.0688 | -106.2332
18.0725 | -106.2269
18.0761 | -106.2198
18.0798 | -106.2134
18.0835 | -106.2099
18.0871 | -106.2092
18.0908 | -106.2069
18.0945 | -106.2066
18.0982 | -106.2062
18.1018 | -106.2073
18.1055 | -106.2093
18.1092 | -106.2104
18.1129 | -106.2135
18.1165 | -106.2185
18.1202 | -106.2258
18.1239 | -106.2324
18.1276 | -106.2395
18.1312 | -106.2488
18.1349 | -106.2575
18.1386 | -106.2653
18.1422 | -106.2712
18.1459 | -106.2753
18.1496 | -106.2827
18.1533 | -106.2904
18.1569 | -106.2986
18.1606 | -106.3097
18.1643 | -106.3193
18.168 | -106.327
18.1716 | -106.3351
18.1753 | -106.3397
18.179 | -106.345
18.1827 | -106.3514
18.1863 | -106.3593
18.19 | -106.3666
18.1937 | -106.3712
18.1973 | -106.3763
18.201 | -106.3836
18.2047 | -106.3898
18.2084 | -106.3971
18.212 | -106.4049
18.2157 | -106.4096
18.2194 | -106.4142
18.2231 | -106.4174
18.2267 | -106.4227
18.2304 | -106.4305
18.2341 | -106.4378
18.2378 | -106.4439
18.2414 | -106.4507
18.2451 | -106.4579
18.2488 | -106.4657
18.2524 | -106.4742
18.2561 | -106.4793
18.2598 | -106.4843
18.2635 | -106.4904
18.2671 | -106.499
18.2708 | -106.5041
18.2745 | -106.5102
18.2782 | -106.5148
18.2818 | -106.5221
18.2855 | -106.5293
18.2892 | -106.5342
18.2929 | -106.5383
18.2965 | -106.5453
18.3002 | -106.5515
18.3039 | -106.5545
18.3075 | -106.5556
18.3112 | -106.5561
18.3149 | -106.5585
18.3186 | -106.5572
18.3222 | -106.5539
18.3259 | -106.5479
18.3296 | -106.5397
18.3333 | -106.5296
18.3369 | -106.5203
18.3406 | -106.5096
18.3443 | -106.4989
18.348 | -106.4881
18.3516 | -106.4783
18.3553 | -106.4697
18.359 | -106.4598
18.3626 | -106.4483
18.3663 | -106.4378
18.37 | -106.4277
18.3737 | -106.4182
18.3773 | -106.409
18.381 | -106.3999
18.3847 | -106.393
18.3884 | -106.3873
18.392 | -106.3809
18.3957 | -106.3766
18.3994 | -106.3724
18.403 | -106.3706
18.4067 | -106.3719
18.4104 | -106.371
18.4141 | -106.3724
18.4177 | -106.3725
18.4214 | -106.3714
18.4251 | -106.375
18.4288 | -106.3801
18.4324 | -106.3862
18.4361 | -106.394
18.4398 | -106.4011
18.4435 | -106.4084
18.4471 | -106.4158
18.4508 | -106.4247
18.4545 | -106.4349
18.4581 | -106.4467
18.4618 | -106.4593
18.4655 | -106.4712
18.4692 | -106.4848
18.4728 | -106.4984
18.4765 | -106.5145
18.4802 | -106.5295
18.4839 | -106.5404
18.4875 | -106.552
18.4912 | -106.5643
18.4949 | -106.5746
18.4986 | -106.5846
18.5022 | -106.5947
18.5059 | -106.6053
18.5096 | -106.6149
18.5132 | -106.623
18.5169 | -106.6327
18.5206 | -106.6426
18.5243 | -106.6505
18.5279 | -106.6542
18.5316 | -106.6553
18.5353 | -106.6566
18.539 | -106.6595
18.5426 | -106.6619
18.5463 | -106.6604
18.55 | -106.6596
18.5537 | -106.6594
18.5573 | -106.6598
18.561 | -106.6583
18.5647 | -106.6551
18.5683 | -106.6538
18.572 | -106.6523
18.5757 | -106.6505
18.5794 | -106.6454
18.583 | -106.6419
18.5867 | -106.6397
18.5904 | -106.6409
18.5941 | -106.6423
18.5977 | -106.6426
18.6014 | -106.6396
18.6051 | -106.6377
18.6088 | -106.6341
18.6124 | -106.6308
18.6161 | -106.6286
18.6198 | -106.6266
18.6234 | -106.6253
18.6271 | -106.6245
18.6308 | -106.6225
18.6345 | -106.621
18.6381 | -106.6184
18.6418 | -106.6149
18.6455 | -106.6103
18.6492 | -106.6069
18.6528 | -106.605
18.6565 | -106.6025
18.6602 | -106.5976
18.6639 | -106.5925
18.6675 | -106.5867
18.6712 | -106.579
18.6749 | -106.5733
18.6785 | -106.5671
18.6822 | -106.5622
18.6859 | -106.5569
18.6896 | -106.5491
18.6932 | -106.5433
18.6969 | -106.5365
18.7006 | -106.5313
18.7043 | -106.5272
18.7079 | -106.5209
18.7116 | -106.515
18.7153 | -106.5115
18.7189 | -106.5073
18.7226 | -106.5023
18.7263 | -106.4977
18.73 | -106.4963
18.7336 | -106.4953
18.7373 | -106.4942
18.741 | -106.4948
18.7447 | -106.4995
18.7483 | -106.5068
18.752 | -106.5103
18.7557 | -106.5143
18.7594 | -106.5207
18.763 | -106.5272
18.7667 | -106.5338
18.7704 | -106.5386
18.774 | -106.5451
18.7777 | -106.5539
18.7814 | -106.5627
18.7851 | -106.5712
18.7887 | -106.5779
18.7924 | -106.5865
18.7961 | -106.5973
18.7998 | -106.6046
18.8034 | -106.61
18.8071 | -106.6152
18.8108 | -106.6196
18.8145 | -106.6254
18.8181 | -106.63
18.8218 | -106.6339
18.8255 | -106.6369
18.8291 | -106.6381
18.8328 | -106.6415
18.8365 | -106.6429
18.8402 | -106.6444
18.8438 | -106.6444
18.8475 | -106.6439
18.8512 | -106.6425
18.8549 | -106.6371
18.8585 | -106.6299
18.8622 | -106.6237
18.8659 | -106.6157
18.8696 | -106.6103
18.8732 | -106.605
18.8769 | -106.5968
18.8806 | -106.5874
18.8842 | -106.5764
18.8879 | -106.5658
18.8916 | -106.555
18.8953 | -106.5439
18.8989 | -106.5329
18.9026 | -106.5202
18.9063 | -106.5074
18.91 | -106.496
18.9136 | -106.4841
18.9173 | -106.4715
18.921 | -106.4589
18.9247 | -106.4476
18.9283 | -106.436
18.932 | -106.4259
18.9357 | -106.418
18.9393 | -106.4102
18.943 | -106.3988
18.9467 | -106.387
18.9504 | -106.3751
18.954 | -106.3637
18.9577 | -106.3538
18.9614 | -106.344
18.9651 | -106.3331
18.9687 | -106.3229
18.9724 | -106.3121
18.9761 | -106.3013
18.9798 | -106.2918
18.9834 | -106.2812
18.9871 | -106.2715
18.9908 | -106.2618
18.9944 | -106.2525
18.9981 | -106.2421
19.0018 | -106.2316
19.0055 | -106.2237
19.0091 | -106.2159
19.0128 | -106.2078
19.0165 | -106.1994
19.0202 | -106.1927
19.0238 | -106.1861
19.0275 | -106.1791
19.0312 | -106.1729
19.0348 | -106.1696
19.0385 | -106.1683
19.0422 | -106.1662
19.0459 | -106.1656
19.0495 | -106.1667
19.0532 | -106.1703
19.0569 | -106.1764
19.0606 | -106.183
19.0642 | -106.191
19.0679 | -106.1993
19.0716 | -106.2091
19.0753 | -106.2229
19.0789 | -106.2314
19.0826 | -106.2415
19.0863 | -106.2516
19.0899 | -106.263
19.0936 | -106.2755
19.0973 | -106.2871
19.101 | -106.299
19.1046 | -106.3087
19.1083 | -106.3171
19.112 | -106.3232
19.1157 | -106.3321
19.1193 | -106.3413
19.123 | -106.3506
19.1267 | -106.3592
19.1304 | -106.3665
19.134 | -106.3733
19.1377 | -106.3812
19.1414 | -106.3857
19.145 | -106.3885
19.1487 | -106.393
19.1524 | -106.3968
19.1561 | -106.3985
19.1597 | -106.3988
19.1634 | -106.4003
19.1671 | -106.4013
19.1708 | -106.4011
19.1744 | -106.4032
19.1781 | -106.4061
19.1818 | -106.4063
19.1855 | -106.4047
19.1891 | -106.4033
19.1928 | -106.4011
19.1965 | -106.4008
19.2001 | -106.4014
19.2038 | -106.4016
19.2075 | -106.4006
19.2112 | -106.4
19.2148 | -106.3982
19.2185 | -106.3986
19.2222 | -106.3998
19.2259 | -106.4018
19.2295 | -106.405
19.2332 | -106.4082
19.2369 | -106.4123
19.2406 | -106.4149
19.2442 | -106.417
19.2479 | -106.4234
19.2516 | -106.4298
19.2552 | -106.4351
19.2589 | -106.4417
19.2626 | -106.4485
19.2663 | -106.4555
19.2699 | -106.461
19.2736 | -106.4681
19.2773 | -106.4752
19.281 | -106.4824
19.2846 | -106.4901
19.2883 | -106.4983
19.292 | -106.5057
19.2957 | -106.5127
19.2993 | -106.52
19.303 | -106.527
19.3067 | -106.5325
19.3103 | -106.5367
19.314 | -106.5417
19.3177 | -106.5451
19.3214 | -106.5455
19.325 | -106.5434
19.3287 | -106.5408
19.3324 | -106.5368
19.3361 | -106.5335
19.3397 | -106.5323
19.3434 | -106.53
19.3471 | -106.5278
19.3508 | -106.5268
19.3544 | -106.5211
19.3581 | -106.5146
19.3618 | -106.5085
19.3654 | -106.5022
19.3691 | -106.4948
19.3728 | -106.4883
19.3765 | -106.4832
19.3801 | -106.4777
19.3838 | -106.4728
19.3875 | -106.4686
19.3912 | -106.4615
19.3948 | -106.4537
19.3985 | -106.4462
19.4022 | -106.4387
19.4059 | -106.4322
19.4095 | -106.4248
19.4132 | -106.4178
19.4169 | -106.4117
19.4205 | -106.4068
19.4242 | -106.4036
19.4279 | -106.4009
19.4316 | -106.3975
19.4352 | -106.3915
19.4389 | -106.3865
19.4426 | -106.3828
19.4463 | -106.38
19.4499 | -106.3762
19.4536 | -106.3719
19.4573 | -106.369
19.4609 | -106.3666
19.4646 | -106.3669
19.4683 | -106.368
19.472 | -106.3647
19.4756 | -106.3614
19.4793 | -106.3595
19.483 | -106.3584
19.4867 | -106.3568
19.4903 | -106.3581
19.494 | -106.358
19.4977 | -106.3578
19.5014 | -106.3554
19.505 | -106.3524
19.5087 | -106.3492
19.5124 | -106.347
19.516 | -106.3474
19.5197 | -106.3494
19.5234 | -106.347
19.5271 | -106.3416
19.5307 | -106.3381
19.5344 | -106.3337
19.5381 | -106.3295
19.5418 | -106.3251
19.5454 | -106.322
19.5491 | -106.3199
19.5528 | -106.316
19.5565 | -106.3113
19.5601 | -106.3062
19.5638 | -106.3009
19.5675 | -106.2946
19.5711 | -106.2881
19.5748 | -106.282
19.5785 | -106.2762
19.5822 | -106.2676
19.5858 | -106.261
19.5895 | -106.2541
19.5932 | -106.2455
19.5969 | -106.2383
19.6005 | -106.2329
19.6042 | -106.2262
19.6079 | -106.2187
19.6116 | -106.2127
19.6152 | -106.2066
19.6189 | -106.1999
19.6226 | -106.1943
19.6262 | -106.1888
19.6299 | -106.1834
19.6336 | -106.1819
19.6373 | -106.181
19.6409 | -106.1779
19.6446 | -106.1744
19.6483 | -106.1709
19.652 | -106.1689
19.6556 | -106.1688
19.6593 | -106.1694
19.663 | -106.1701
19.6667 | -106.1732
19.6703 | -106.177
19.674 | -106.1831
19.6777 | -106.1883
19.6813 | -106.1926
19.685 | -106.1971
19.6887 | -106.2033
19.6924 | -106.2096
19.696 | -106.2175
19.6997 | -106.2267
19.7034 | -106.2348
19.7071 | -106.2432
19.7107 | -106.2542
19.7144 | -106.2657
19.7181 | -106.2753
19.7218 | -106.2856
19.7254 | -106.2955
19.7291 | -106.303
19.7328 | -106.311
19.7364 | -106.3182
19.7401 | -106.3253
19.7438 | -106.3281
19.7475 | -106.3321
19.7511 | -106.3393
19.7548 | -106.3474
19.7585 | -106.3524
19.7622 | -106.3553
19.7658 | -106.3583
19.7695 | -106.362
19.7732 | -106.3633
19.7768 | -106.3642
19.7805 | -106.3674
19.7842 | -106.3666
19.7879 | -106.365
19.7915 | -106.3629
19.7952 | -106.359
19.7989 | -106.357
19.8026 | -106.3561
19.8062 | -106.3551
19.8099 | -106.3515
19.8136 | -106.3472
19.8173 | -106.3418
19.8209 | -106.3363
19.8246 | -106.3308
19.8283 | -106.3232
19.8319 | -106.3157
19.8356 | -106.3082
19.8393 | -106.3007
19.843 | -106.2944
19.8466 | -106.2908
19.8503 | -106.2861
19.854 | -106.2788
19.8577 | -106.2702
19.8613 | -106.2652
19.865 | -106.259
19.8687 | -106.2526
19.8724 | -106.2503
19.876 | -106.2479
19.8797 | -106.2438
19.8834 | -106.2406
19.887 | -106.2367
19.8907 | -106.2359
19.8944 | -106.2352
19.8981 | -106.2328
19.9017 | -106.2302
19.9054 | -106.2283
19.9091 | -106.2295
19.9128 | -106.2306
19.9164 | -106.2316
19.9201 | -106.2337
19.9238 | -106.2394
19.9275 | -106.2448
19.9311 | -106.2497
19.9348 | -106.2541
19.9385 | -106.2582
19.9421 | -106.2645
19.9458 | -106.2711
19.9495 | -106.2795
19.9532 | -106.2882
19.9568 | -106.2952
19.9605 | -106.3012
19.9642 | -106.3116
19.9679 | -106.3229
19.9715 | -106.3326
19.9752 | -106.3427
19.9789 | -106.35
19.9826 | -106.3575
19.9862 | -106.3663
19.9899 | -106.3776
19.9936 | -106.3873
19.9972 | -106.3984
20.0009 | -106.4112
20.0046 | -106.4211
20.0083 | -106.4328
20.0119 | -106.4458
20.0156 | -106.4567
20.0193 | -106.4703
20.023 | -106.4832
20.0266 | -106.4944
20.0303 | -106.5037
20.034 | -106.5128
20.0377 | -106.5188
20.0413 | -106.5249
20.045 | -106.5326
20.0487 | -106.5412
20.0523 | -106.5459
20.056 | -106.5507
20.0597 | -106.5559
20.0634 | -106.5598
20.067 | -106.5651
20.0707 | -106.569
20.0744 | -106.5728
20.0781 | -106.5748
20.0817 | -106.5727
20.0854 | -106.5725
20.0891 | -106.5716
20.0927 | -106.5712
20.0964 | -106.569
20.1001 | -106.567
20.1038 | -106.5639
20.1074 | -106.5594
20.1111 | -106.5558
20.1148 | -106.5531
20.1185 | -106.546
20.1221 | -106.5385
20.1258 | -106.532
20.1295 | -106.5238
20.1332 | -106.5147
20.1368 | -106.5052
20.1405 | -106.4993
20.1442 | -106.4939
20.1478 | -106.4883
20.1515 | -106.4816
20.1552 | -106.475
20.1589 | -106.4663
20.1625 | -106.4571
20.1662 | -106.448
20.1699 | -106.4399
20.1736 | -106.4338
20.1772 | -106.4268
20.1809 | -106.419
20.1846 | -106.4111
20.1883 | -106.404
20.1919 | -106.3966
20.1956 | -106.3876
20.1993 | -106.3803
20.2029 | -106.3722
20.2066 | -106.3647
20.2103 | -106.3562
20.214 | -106.3485
20.2176 | -106.3423
20.2213 | -106.3341
20.225 | -106.3238
20.2287 | -106.3148
20.2323 | -106.3067
20.236 | -106.2941
20.2397 | -106.284
20.2434 | -106.2743
20.247 | -106.2656
20.2507 | -106.2562
20.2544 | -106.2473
20.258 | -106.2373
20.2617 | -106.2285
20.2654 | -106.2213
20.2691 | -106.2138
20.2727 | -106.2061
20.2764 | -106.1967
20.2801 | -106.1861
20.2838 | -106.1761
20.2874 | -106.1663
20.2911 | -106.1583
20.2948 | -106.1486
20.2985 | -106.1398
20.3021 | -106.1316
20.3058 | -106.1222
20.3095 | -106.1151
20.3131 | -106.1108
20.3168 | -106.1052
20.3205 | -106.0982
20.3242 | -106.0907
20.3278 | -106.0861
20.3315 | -106.0799
20.3352 | -106.072
20.3389 | -106.0665
20.3425 | -106.0643
20.3462 | -106.062
20.3499 | -106.0599
20.3536 | -106.0581
20.3572 | -106.057
20.3609 | -106.0546
20.3646 | -106.0539
20.3682 | -106.0551
20.3719 | -106.0553
20.3756 | -106.057
20.3793 | -106.0589
20.3829 | -106.0614
20.3866 | -106.0621
20.3903 | -106.0654
20.394 | -106.0697
20.3976 | -106.0731
20.4013 | -106.078
20.405 | -106.0847
20.4087 | -106.088
20.4123 | -106.0885
20.416 | -106.0906
20.4197 | -106.0946
20.4233 | -106.0988
20.427 | -106.0984
20.4307 | -106.0987
20.4344 | -106.0979
20.438 | -106.0946
20.4417 | -106.0925
20.4454 | -106.0904
20.4491 | -106.0881
20.4527 | -106.0863
20.4564 | -106.0831
20.4601 | -106.0793
20.4637 | -106.0732
20.4674 | -106.0661
20.4711 | -106.0603
20.4748 | -106.0531
20.4784 | -106.0416
20.4821 | -106.033
20.4858 | -106.0225
20.4895 | -106.0107
20.4931 | -105.9998
20.4968 | -105.9871
20.5005 | -105.9743
20.5042 | -105.9608
20.5078 | -105.948
20.5115 | -105.9363
20.5152 | -105.9249
20.5188 | -105.9121
20.5225 | -105.8981
20.5262 | -105.8874
20.5299 | -105.8742
20.5335 | -105.8626
20.5372 | -105.8508
20.5409 | -105.8378
20.5446 | -105.8239
20.5482 | -105.8122
20.5519 | -105.8001
20.5556 | -105.79
20.5593 | -105.7791
20.5629 | -105.7699
20.5666 | -105.7621
20.5703 | -105.758
20.5739 | -105.755
20.5776 | -105.7505
20.5813 | -105.7463
20.585 | -105.7437
20.5886 | -105.7413
20.5923 | -105.7402
20.596 | -105.7401
20.5997 | -105.74
20.6033 | -105.7407
20.607 | -105.7425
20.6107 | -105.7462
20.6144 | -105.7532
20.618 | -105.7602
20.6217 | -105.7681
20.6254 | -105.7764
20.629 | -105.7856
20.6327 | -105.792
20.6364 | -105.7982
20.6401 | -105.8057
20.6437 | -105.8147
20.6474 | -105.8257
20.6511 | -105.8369
20.6548 | -105.8465
20.6584 | -105.8571
20.6621 | -105.868
20.6658 | -105.8779
20.6695 | -105.8852
20.6731 | -105.8949
20.6768 | -105.9048
20.6805 | -105.9153
20.6841 | -105.9224
20.6878 | -105.9279
20.6915 | -105.9359
20.6952 | -105.9441
20.6988 | -105.9528
20.7025 | -105.9633
20.7062 | -105.9723
20.7099 | -105.9793
20.7135 | -105.985
20.7172 | -105.9894
20.7209 | -105.9926
20.7246 | -105.9952
20.7282 | -105.9982
20.7319 | -106.0011
20.7356 | -106.0028
20.7392 | -106.0027
20.7429 | -106.0015
20.7466 | -106.0002
20.7503 | -105.9977
20.7539 | -105.9934
20.7576 | -105.9909
20.7613 | -105.9882
20.765 | -105.983
20.7686 | -105.9772
20.7723 | -105.9733
20.776 | -105.9664
20.7797 | -105.9601
20.7833 | -105.9531
20.787 | -105.9448
20.7907 | -105.9358
20.7943 | -105.9268
20.798 | -105.918
20.8017 | -105.9077
20.8054 | -105.8979
20.809 | -105.889
20.8127 | -105.8796
20.8164 | -105.8715
20.8201 | -105.8621
20.8237 | -105.8542
20.8274 | -105.8467
20.8311 | -105.8402
20.8347 | -105.8341
20.8384 | -105.8307
20.8421 | -105.8274
20.8458 | -105.8251
20.8494 | -105.8236
20.8531 | -105.8221
20.8568 | -105.819
20.8605 | -105.8172
20.8641 | -105.8184
20.8678 | -105.8183
20.8715 | -105.8207
20.8752 | -105.8244
20.8788 | -105.8297
20.8825 | -105.8357
20.8862 | -105.8426
20.8898 | -105.8516
20.8935 | -105.86
20.8972 | -105.8682
20.9009 | -105.878
20.9045 | -105.8876
20.9082 | -105.8951
20.9119 | -105.9036
20.9156 | -105.9117
20.9192 | -105.9205
20.9229 | -105.929
20.9266 | -105.9389
20.9303 | -105.95
20.9339 | -105.9594
20.9376 | -105.9679
20.9413 | -105.9747
20.9449 | -105.9823
20.9486 | -105.9884
20.9523 | -105.9916
20.956 | -105.9944
20.9596 | -105.997
20.9633 | -106.0002
20.967 | -106.0033
20.9707 | -106.0034
20.9743 | -106.0023
20.978 | -106.0002
20.9817 | -105.9987
20.9854 | -105.9973
20.989 | -105.9925
20.9927 | -105.987
20.9964 | -105.9809
21 | -105.973
21.0037 | -105.9645
21.0074 | -105.9566
21.0111 | -105.9488
21.0147 | -105.941
21.0184 | -105.9355
21.0221 | -105.9289
21.0258 | -105.9219
21.0294 | -105.9148
21.0331 | -105.9082
21.0368 | -105.8993
21.0405 | -105.8928
21.0441 | -105.8858
21.0478 | -105.8797
21.0515 | -105.8753
21.0551 | -105.8693
21.0588 | -105.8646
21.0625 | -105.8604
21.0662 | -105.8573
21.0698 | -105.8562
21.0735 | -105.8532
21.0772 | -105.8504
21.0809 | -105.8495
21.0845 | -105.8502
21.0882 | -105.8519
21.0919 | -105.8519
21.0956 | -105.8527
21.0992 | -105.8545
21.1029 | -105.8575
21.1066 | -105.8608
21.1102 | -105.8634
21.1139 | -105.8676
21.1176 | -105.8722
21.1213 | -105.8779
21.1249 | -105.8821
21.1286 | -105.8861
21.1323 | -105.8907
21.136 | -105.8957
21.1396 | -105.9013
21.1433 | -105.9061
21.147 | -105.9097
21.1506 | -105.9167
21.1543 | -105.9253
21.158 | -105.9301
21.1617 | -105.9342
21.1653 | -105.9403
21.169 | -105.9457
21.1727 | -105.95
21.1764 | -105.9527
21.18 | -105.9575
21.1837 | -105.963
21.1874 | -105.9663
21.1911 | -105.9697
21.1947 | -105.9756
21.1984 | -105.9802
21.2021 | -105.9849
21.2057 | -105.9889
21.2094 | -105.9938
21.2131 | -105.9973
21.2168 | -106.002
21.2204 | -106.0076
21.2241 | -106.0132
21.2278 | -106.0186
21.2315 | -106.0233
21.2351 | -106.03
21.2388 | -106.0363
21.2425 | -106.0427
21.2462 | -106.0497
21.2498 | -106.0583
21.2535 | -106.0674
21.2572 | -106.0753
21.2608 | -106.084
21.2645 | -106.0942
21.2682 | -106.1044
21.2719 | -106.1133
21.2755 | -106.1231
21.2792 | -106.1324
21.2829 | -106.1418
21.2866 | -106.1522
21.2902 | -106.1626
21.2939 | -106.1714
21.2976 | -106.1808
21.3013 | -106.1905
21.3049 | -106.2022
21.3086 | -106.2116
21.3123 | -106.2203
21.3159 | -106.2292
21.3196 | -106.2358
21.3233 | -106.2422
21.327 | -106.2489
21.3306 | -106.2555
21.3343 | -106.2604
21.338 | -106.2676
21.3417 | -106.2707
21.3453 | -106.2745
21.349 | -106.2786
21.3527 | -106.2811
21.3564 | -106.2822
21.36 | -106.2823
21.3637 | -106.282
21.3674 | -106.2816
21.371 | -106.2791
21.3747 | -106.2769
21.3784 | -106.273
21.3821 | -106.2683
21.3857 | -106.2618
21.3894 | -106.2526
21.3931 | -106.2451
21.3968 | -106.2348
21.4004 | -106.2235
21.4041 | -106.2117
21.4078 | -106.1991
21.4115 | -106.1884
21.4151 | -106.1762
21.4188 | -106.1631
21.4225 | -106.1508
21.4261 | -106.1399
21.4298 | -106.1275
21.4335 | -106.1146
21.4372 | -106.1012
21.4408 | -106.0883
21.4445 | -106.0752
21.4482 | -106.062
21.4519 | -106.0495
21.4555 | -106.0383
21.4592 | -106.0303
21.4629 | -106.0187
21.4666 | -106.0084
21.4702 | -105.9989
21.4739 | -105.9899
21.4776 | -105.9823
21.4812 | -105.9752
21.4849 | -105.9692
21.4886 | -105.9652
21.4923 | -105.9617
21.4959 | -105.9587
21.4996 | -105.9538
21.5033 | -105.9523
21.507 | -105.9522
21.5106 | -105.9526
21.5143 | -105.9542
21.518 | -105.9554
21.5216 | -105.9559
21.5253 | -105.9557
21.529 | -105.9567
21.5327 | -105.9593
21.5363 | -105.9633
21.54 | -105.9674
21.5437 | -105.9725
21.5474 | -105.9778
21.551 | -105.9841
21.5547 | -105.9916
21.5584 | -105.9987
21.5621 | -106.0077
21.5657 | -106.014
21.5694 | -106.0203
21.5731 | -106.0267
21.5767 | -106.0327
21.5804 | -106.0401
21.5841 | -106.05
21.5878 | -106.0587
21.5914 | -106.0656
21.5951 | -106.072
21.5988 | -106.0792
21.6025 | -106.0882
21.6061 | -106.0976
21.6098 | -106.1053
21.6135 | -106.1127
21.6172 | -106.1185
21.6208 | -106.1234
21.6245 | -106.1283
21.6282 | -106.1317
21.6318 | -106.1346
21.6355 | -106.1373
21.6392 | -106.1388
21.6429 | -106.1409
21.6465 | -106.1429
21.6502 | -106.1424
21.6539 | -106.1391
21.6576 | -106.1342
21.6612 | -106.129
21.6649 | -106.1239
21.6686 | -106.121
21.6723 | -106.1166
21.6759 | -106.1128
21.6796 | -106.1086
21.6833 | -106.1038
21.6869 | -106.0976
21.6906 | -106.0906
21.6943 | -106.0835
21.698 | -106.076
21.7016 | -106.0694
21.7053 | -106.0619
21.709 | -106.0509
21.7127 | -106.0413
21.7163 | -106.0294
21.72 | -106.0162
21.7237 | -106.0037
21.7274 | -105.9879
21.731 | -105.9749
21.7347 | -105.9633
21.7384 | -105.9525
21.742 | -105.9406
21.7457 | -105.9279
21.7494 | -105.9165
21.7531 | -105.9079
21.7567 | -105.8979
21.7604 | -105.887
21.7641 | -105.8767
21.7678 | -105.8659
21.7714 | -105.855
21.7751 | -105.8441
21.7788 | -105.8344
21.7825 | -105.8251
21.7861 | -105.8178
21.7898 | -105.8092
21.7935 | -105.7983
21.7971 | -105.788
21.8008 | -105.7807
21.8045 | -105.7741
21.8082 | -105.7677
21.8118 | -105.7614
21.8155 | -105.755
21.8192 | -105.7502
21.8229 | -105.7458
21.8265 | -105.7407
21.8302 | -105.7373
21.8339 | -105.7352
21.8376 | -105.7326
21.8412 | -105.7297
21.8449 | -105.7245
21.8486 | -105.72
21.8522 | -105.7176
21.8559 | -105.7154
21.8596 | -105.7131
21.8633 | -105.7105
21.8669 | -105.7087
21.8706 | -105.7052
21.8743 | -105.704
21.878 | -105.7025
21.8816 | -105.7028
21.8853 | -105.7057
21.889 | -105.7065
21.8926 | -105.7046
21.8963 | -105.7054
21.9 | -105.7079
21.9037 | -105.7103
21.9073 | -105.713
21.911 | -105.7132
21.9147 | -105.713
21.9184 | -105.7135
21.922 | -105.7143
21.9257 | -105.7162
21.9294 | -105.7187
21.9331 | -105.7199
21.9367 | -105.7219
21.9404 | -105.7234
21.9441 | -105.7251
21.9477 | -105.7258
21.9514 | -105.7257
21.9551 | -105.7252
21.9588 | -105.7249
21.9624 | -105.7265
21.9661 | -105.7274
21.9698 | -105.7278
21.9735 | -105.728
21.9771 | -105.7271
21.9808 | -105.727
21.9845 | -105.7262
21.9882 | -105.7252
21.9918 | -105.7245
21.9955 | -105.7227
21.9992 | -105.7225
22.0028 | -105.7228
22.0065 | -105.7213
22.0102 | -105.7194
22.0139 | -105.7189
22.0175 | -105.7183
22.0212 | -105.7172
22.0249 | -105.7177
22.0286 | -105.7195
22.0322 | -105.7194
22.0359 | -105.7205
22.0396 | -105.7208
22.0433 | -105.7205
22.0469 | -105.7179
22.0506 | -105.7165
22.0543 | -105.7156
22.0579 | -105.7153
22.0616 | -105.7139
22.0653 | -105.711
22.069 | -105.7096
22.0726 | -105.7078
22.0763 | -105.7044
22.08 | -105.7028
22.0837 | -105.7019
22.0873 | -105.6987
22.091 | -105.6944
22.0947 | -105.6921
22.0984 | -105.6901
22.102 | -105.6865
22.1057 | -105.6832
22.1094 | -105.6798
22.113 | -105.6741
22.1167 | -105.6682
22.1204 | -105.6648
22.1241 | -105.6606
22.1277 | -105.6563
22.1314 | -105.6514
22.1351 | -105.6469
22.1388 | -105.6432
22.1424 | -105.6383
22.1461 | -105.6329
22.1498 | -105.6277
22.1535 | -105.6246
22.1571 | -105.6234
22.1608 | -105.6216
22.1645 | -105.6178
22.1681 | -105.6151
22.1718 | -105.6133
22.1755 | -105.612
22.1792 | -105.6122
22.1828 | -105.6141
22.1865 | -105.6148
22.1902 | -105.615
22.1939 | -105.616
22.1975 | -105.6166
22.2012 | -105.6166
22.2049 | -105.6201
22.2085 | -105.6237
22.2122 | -105.6274
22.2159 | -105.6327
22.2196 | -105.6385
22.2232 | -105.6447
22.2269 | -105.6519
22.2306 | -105.6608
22.2343 | -105.6706
22.2379 | -105.6814
22.2416 | -105.6913
22.2453 | -105.7035
22.249 | -105.7152
22.2526 | -105.7265
22.2563 | -105.7445
22.26 | -105.7563
22.2636 | -105.769
22.2673 | -105.7801
22.271 | -105.7937
22.2747 | -105.8084
22.2783 | -105.8227
22.282 | -105.8356
22.2857 | -105.8486
22.2894 | -105.8614
22.293 | -105.8746
22.2967 | -105.8866
22.3004 | -105.8982
22.3041 | -105.9097
22.3077 | -105.9208
22.3114 | -105.9331
22.3151 | -105.9453
22.3187 | -105.9564
22.3224 | -105.9677
22.3261 | -105.978
22.3298 | -105.9859
22.3334 | -105.9942
22.3371 | -106.001
22.3408 | -106.0065
22.3445 | -106.01
22.3481 | -106.0163
22.3518 | -106.0212
22.3555 | -106.0232
22.3592 | -106.0251
22.3628 | -106.0271
22.3665 | -106.026
22.3702 | -106.0237
22.3738 | -106.0229
22.3775 | -106.021
22.3812 | -106.0186
22.3849 | -106.0137
22.3885 | -106.005
22.3922 | -105.9968
22.3959 | -105.9846
22.3996 | -105.9735
22.4032 | -105.9588
22.4069 | -105.9414
22.4106 | -105.9207
22.4143 | -105.8965
22.4179 | -105.8732
22.4216 | -105.8548
22.4253 | -105.8365
22.4289 | -105.8157
22.4326 | -105.7926
22.4363 | -105.7674
22.44 | -105.7399
22.4436 | -105.7122
22.4473 | -105.6854
22.451 | -105.6563
22.4547 | -105.6259
22.4583 | -105.5954
22.462 | -105.5651
22.4657 | -105.5317
22.4694 | -105.4961
22.473 | -105.4616
22.4767 | -105.4275
22.4804 | -105.394
22.484 | -105.3594
22.4877 | -105.3233
22.4914 | -105.2878
22.4951 | -105.2527
22.4987 | -105.2185
22.5024 | -105.185
22.5061 | -105.1528
22.5098 | -105.1203
22.5134 | -105.0878
22.5171 | -105.059
22.5208 | -105.0302
22.5245 | -105.0011
22.5281 | -104.974
22.5318 | -104.9472
22.5355 | -104.9239
22.5391 | -104.9034
22.5428 | -104.8843
22.5465 | -104.8661
22.5502 | -104.8496
22.5538 | -104.8345
22.5575 | -104.8233
22.5612 | -104.8145
22.5649 | -104.8077
22.5685 | -104.8033
22.5722 | -104.8004
22.5759 | -104.7994
22.5795 | -104.8009
22.5832 | -104.8041
22.5869 | -104.8102
22.5906 | -104.8172
22.5942 | -104.8253
22.5979 | -104.8353
22.6016 | -104.8496
22.6053 | -104.8654
22.6089 | -104.8804
22.6126 | -104.8957
22.6163 | -104.9128
22.62 | -104.9316
22.6236 | -104.9518
22.6273 | -104.9727
22.631 | -104.9935
22.6346 | -105.0129
22.6383 | -105.034
22.642 | -105.0551
22.6457 | -105.0763
22.6493 | -105.0961
22.653 | -105.1164
22.6567 | -105.1351
22.6604 | -105.1546
22.664 | -105.1741
22.6677 | -105.1906
22.6714 | -105.2074
22.6751 | -105.225
22.6787 | -105.241
22.6824 | -105.2569
22.6861 | -105.2732
22.6897 | -105.2906
22.6934 | -105.3045
22.6971 | -105.3182
22.7008 | -105.331
22.7044 | -105.3451
22.7081 | -105.3593
22.7118 | -105.3735
22.7155 | -105.39
22.7191 | -105.4036
22.7228 | -105.4181
22.7265 | -105.4323
22.7302 | -105.4468
22.7338 | -105.4634
22.7375 | -105.479
22.7412 | -105.4885
22.7448 | -105.4968
22.7485 | -105.5074
22.7522 | -105.5193
22.7559 | -105.5322
22.7595 | -105.5456
22.7632 | -105.5569
22.7669 | -105.5665
22.7706 | -105.5791
22.7742 | -105.5916
22.7779 | -105.6035
22.7816 | -105.6142
22.7853 | -105.6264
22.7889 | -105.6393
22.7926 | -105.6501
22.7963 | -105.6614
22.7999 | -105.6725
22.8036 | -105.6843
22.8073 | -105.6959
22.811 | -105.7065
22.8146 | -105.7165
22.8183 | -105.7259
22.822 | -105.7346
22.8257 | -105.7441
22.8293 | -105.7522
22.833 | -105.7609
22.8367 | -105.7682
22.8404 | -105.7739
22.844 | -105.778
22.8477 | -105.7811
22.8514 | -105.7845
22.855 | -105.7893
22.8587 | -105.7925
22.8624 | -105.7956
22.8661 | -105.7984
22.8697 | -105.7999
22.8734 | -105.8
22.8771 | -105.8004
22.8808 | -105.8006
22.8844 | -105.7991
22.8881 | -105.7969
22.8918 | -105.7946
22.8955 | -105.7902
22.8991 | -105.7868
22.9028 | -105.7838
22.9065 | -105.7805
22.9101 | -105.7752
22.9138 | -105.7665
22.9175 | -105.7596
22.9212 | -105.7522
22.9248 | -105.7424
22.9285 | -105.7312
22.9322 | -105.7222
22.9359 | -105.7134
22.9395 | -105.7043
22.9432 | -105.6955
22.9469 | -105.685
22.9505 | -105.6735
22.9542 | -105.6637
22.9579 | -105.6556
22.9616 | -105.6469
22.9652 | -105.6355
22.9689 | -105.6261
22.9726 | -105.6182
22.9763 | -105.6091
22.9799 | -105.6013
22.9836 | -105.5924
22.9873 | -105.584
22.991 | -105.5744
22.9946 | -105.5645
22.9983 | -105.5555
23.002 | -105.5472
23.0056 | -105.5402
23.0093 | -105.5337
23.013 | -105.525
23.0167 | -105.5164
23.0203 | -105.5083
23.024 | -105.5003
23.0277 | -105.4927
23.0314 | -105.4861
23.035 | -105.4782
23.0387 | -105.4696
23.0424 | -105.4636
23.0461 | -105.4592
23.0497 | -105.4555
23.0534 | -105.4503
23.0571 | -105.4415
23.0607 | -105.4357
23.0644 | -105.4315
23.0681 | -105.4252
23.0718 | -105.4221
23.0754 | -105.4204
23.0791 | -105.417
23.0828 | -105.4154
23.0865 | -105.4156
23.0901 | -105.4157
23.0938 | -105.4137
23.0975 | -105.414
23.1012 | -105.4149
23.1048 | -105.4158
23.1085 | -105.4171
23.1122 | -105.4182
23.1158 | -105.4188
23.1195 | -105.422
23.1232 | -105.4254
23.1269 | -105.4278
23.1305 | -105.4301
23.1342 | -105.433
23.1379 | -105.4346
23.1416 | -105.4368
23.1452 | -105.4408
23.1489 | -105.4441
23.1526 | -105.4471
23.1563 | -105.4505
23.1599 | -105.454
23.1636 | -105.4589
23.1673 | -105.4619
23.1709 | -105.4647
23.1746 | -105.4693
23.1783 | -105.4712
23.182 | -105.4734
23.1856 | -105.4747
23.1893 | -105.4756
23.193 | -105.476
23.1967 | -105.4754
23.2003 | -105.4767
23.204 | -105.4775
23.2077 | -105.4757
23.2113 | -105.4748
23.215 | -105.4755
23.2187 | -105.475
23.2224 | -105.474
23.226 | -105.4711
23.2297 | -105.4685
23.2334 | -105.4676
23.2371 | -105.4668
23.2407 | -105.4669
23.2444 | -105.4656
23.2481 | -105.464
23.2518 | -105.4619
23.2554 | -105.4614
23.2591 | -105.459
23.2628 | -105.4543
23.2664 | -105.4514
23.2701 | -105.4468
23.2738 | -105.4436
23.2775 | -105.4431
23.2811 | -105.4411
23.2848 | -105.4383
23.2885 | -105.4347
23.2922 | -105.433
23.2958 | -105.4296
23.2995 | -105.426
23.3032 | -105.4243
23.3069 | -105.4215
23.3105 | -105.4176
23.3142 | -105.4144
23.3179 | -105.4109
23.3215 | -105.41
23.3252 | -105.4094
23.3289 | -105.408
23.3326 | -105.4057
23.3362 | -105.402
23.3399 | -105.3983
23.3436 | -105.3973
23.3473 | -105.3945
23.3509 | -105.3903
23.3546 | -105.3876
23.3583 | -105.3856
23.362 | -105.3838
23.3656 | -105.3797
23.3693 | -105.3764
23.373 | -105.3738
23.3766 | -105.3694
23.3803 | -105.3652
23.384 | -105.3615
23.3877 | -105.3608
23.3913 | -105.3585
23.395 | -105.3567
23.3987 | -105.3571
23.4024 | -105.3552
23.406 | -105.3525
23.4097 | -105.3506
23.4134 | -105.3474
23.4171 | -105.3447
23.4207 | -105.3428
23.4244 | -105.3416
23.4281 | -105.3401
23.4317 | -105.3392
23.4354 | -105.3385
23.4391 | -105.3378
23.4428 | -105.3375
23.4464 | -105.3378
23.4501 | -105.3378
23.4538 | -105.3373
23.4575 | -105.3373
23.4611 | -105.338
23.4648 | -105.3394
23.4685 | -105.3418
23.4722 | -105.3439
23.4758 | -105.3459
23.4795 | -105.3485
23.4832 | -105.3512
23.4868 | -105.3536
23.4905 | -105.355
23.4942 | -105.3577
23.4979 | -105.3617
23.5015 | -105.3623
23.5052 | -105.3623
23.5089 | -105.3648
23.5126 | -105.3675
23.5162 | -105.3704
23.5199 | -105.3741
23.5236 | -105.378
23.5273 | -105.3829
23.5309 | -105.3869
23.5346 | -105.3909
23.5383 | -105.3956
23.5419 | -105.3999
23.5456 | -105.4028
23.5493 | -105.406
23.553 | -105.4103
23.5566 | -105.4147
23.5603 | -105.4175
23.564 | -105.42
23.5677 | -105.4241
23.5713 | -105.4296
23.575 | -105.4333
23.5787 | -105.4356
23.5824 | -105.4399
23.586 | -105.4434
23.5897 | -105.4466
23.5934 | -105.4527
23.597 | -105.4563
23.6007 | -105.4613
23.6044 | -105.4668
23.6081 | -105.4683
23.6117 | -105.47
23.6154 | -105.4716
23.6191 | -105.4735
23.6228 | -105.4744
23.6264 | -105.4774
23.6301 | -105.4809
23.6338 | -105.483
23.6374 | -105.4854
23.6411 | -105.4889
23.6448 | -105.4924
23.6485 | -105.4941
23.6521 | -105.4956
23.6558 | -105.4968
23.6595 | -105.4966
23.6632 | -105.495
23.6668 | -105.4948
23.6705 | -105.4958
23.6742 | -105.498
23.6779 | -105.4997
23.6815 | -105.5027
23.6852 | -105.5057
23.6889 | -105.5097
23.6925 | -105.5137
23.6962 | -105.5169
23.6999 | -105.5218
23.7036 | -105.5256
23.7072 | -105.5283
23.7109 | -105.5332
23.7146 | -105.5372
23.7183 | -105.5412
23.7219 | -105.5463
23.7256 | -105.5514
23.7293 | -105.558
23.733 | -105.5641
23.7366 | -105.5701
23.7403 | -105.5759
23.744 | -105.584
23.7476 | -105.5923
23.7513 | -105.6003
23.755 | -105.6095
23.7587 | -105.6174
23.7623 | -105.6258
23.766 | -105.6345
23.7697 | -105.6445
23.7734 | -105.6546
23.777 | -105.6644
23.7807 | -105.6746
23.7844 | -105.6862
23.7881 | -105.6957
23.7917 | -105.7057
23.7954 | -105.7156
23.7991 | -105.7243
23.8027 | -105.7333
23.8064 | -105.7414
23.8101 | -105.7489
23.8138 | -105.7563
23.8174 | -105.763
23.8211 | -105.7703
23.8248 | -105.7787
23.8285 | -105.7858
23.8321 | -105.7916
23.8358 | -105.8006
23.8395 | -105.8091
23.8432 | -105.8142
23.8468 | -105.8119
23.8505 | -105.8148
23.8542 | -105.8153
23.8578 | -105.8151
23.8615 | -105.8162
23.8652 | -105.8172
23.8689 | -105.8173
23.8725 | -105.8165
23.8762 | -105.8144
23.8799 | -105.8129
23.8836 | -105.8124
23.8872 | -105.8118
23.8909 | -105.8094
23.8946 | -105.8069
23.8983 | -105.805
23.9019 | -105.8025
23.9056 | -105.798
23.9093 | -105.7918
23.9129 | -105.7878
23.9166 | -105.7845
23.9203 | -105.7802
23.924 | -105.7765
23.9276 | -105.7723
23.9313 | -105.7666
23.935 | -105.7617
23.9387 | -105.7569
23.9423 | -105.7513
23.946 | -105.7491
23.9497 | -105.746
23.9534 | -105.7432
23.957 | -105.7412
23.9607 | -105.7413
23.9644 | -105.7404
23.968 | -105.7383
23.9717 | -105.737
23.9754 | -105.7361
23.9791 | -105.7359
23.9827 | -105.7349
23.9864 | -105.7353
23.9901 | -105.7366
23.9938 | -105.7372
23.9974 | -105.739
24.0011 | -105.7421
24.0048 | -105.7439
24.0084 | -105.7444
24.0121 | -105.7457
24.0158 | -105.7458
24.0195 | -105.7476
24.0231 | -105.7515
24.0268 | -105.755
24.0305 | -105.7575
24.0342 | -105.7593
24.0378 | -105.7624
24.0415 | -105.7671
24.0452 | -105.7718
24.0489 | -105.7741
24.0525 | -105.7785
24.0562 | -105.7839
24.0599 | -105.7864
24.0635 | -105.7894
24.0672 | -105.7906
24.0709 | -105.794
24.0746 | -105.7994
24.0782 | -105.804
24.0819 | -105.8075
24.0856 | -105.8109
24.0893 | -105.8143
24.0929 | -105.8171
24.0966 | -105.8211
24.1003 | -105.8239
24.104 | -105.8248
24.1076 | -105.8248
24.1113 | -105.8241
24.115 | -105.8234
24.1186 | -105.8231
24.1223 | -105.8202
24.126 | -105.8188
24.1297 | -105.8176
24.1333 | -105.8147
24.137 | -105.8121
24.1407 | -105.8108
24.1444 | -105.8092
24.148 | -105.8066
24.1517 | -105.8021
24.1554 | -105.7986
24.1591 | -105.7934
24.1627 | -105.7876
24.1664 | -105.782
24.1701 | -105.7769
24.1737 | -105.7707
24.1774 | -105.7633
24.1811 | -105.7579
24.1848 | -105.7533
24.1884 | -105.7452
24.1921 | -105.7373
24.1958 | -105.7318
24.1995 | -105.7262
24.2031 | -105.7194
24.2068 | -105.7122
24.2105 | -105.7053
24.2142 | -105.6977
24.2178 | -105.6914
24.2215 | -105.6845
24.2252 | -105.6752
24.2288 | -105.668
24.2325 | -105.6611
24.2362 | -105.6518
24.2399 | -105.6409
24.2435 | -105.629
24.2472 | -105.6138
24.2509 | -105.593
24.2546 | -105.5677
24.2582 | -105.5405
24.2619 | -105.4969
24.2656 | -105.4291
24.2692 | -105.3466
24.2729 | -105.2641
24.2766 | -105.1955
24.2803 | -105.1427
24.2839 | -105.0911
24.2876 | -105.0367
24.2913 | -104.9784
24.295 | -104.9172
24.2986 | -104.8528
24.3023 | -104.785
24.306 | -104.7137
24.3097 | -104.6406
24.3133 | -104.5651
24.317 | -104.4865
24.3207 | -104.4041
24.3243 | -104.3224
24.328 | -104.2388
24.3317 | -104.1526
24.3354 | -104.0655
24.339 | -103.9768
24.3427 | -103.8883
24.3464 | -103.7994
24.3501 | -103.7101
24.3537 | -103.6206
24.3574 | -103.5321
24.3611 | -103.4439
24.3648 | -103.3564
24.3684 | -103.2707
24.3721 | -103.1884
24.3758 | -103.1217
24.3794 | -103.0482
24.3831 | -102.9776
24.3868 | -102.9099
24.3905 | -102.8449
24.3941 | -102.785
24.3978 | -102.7295
24.4015 | -102.6774
24.4052 | -102.6296
24.4088 | -102.5862
24.4125 | -102.5478
24.4162 | -102.5148
24.4199 | -102.4861
24.4235 | -102.4632
24.4272 | -102.4456
24.4309 | -102.4339
24.4345 | -102.4263
24.4382 | -102.4239
24.4419 | -102.427
24.4456 | -102.435
24.4492 | -102.4486
24.4529 | -102.4671
24.4566 | -102.4911
24.4603 | -102.5207
24.4639 | -102.5549
24.4676 | -102.5936
24.4713 | -102.6359
24.475 | -102.6827
24.4786 | -102.734
24.4823 | -102.7893
24.486 | -102.8479
24.4896 | -102.9091
24.4933 | -102.9734
24.497 | -103.0415
24.5007 | -103.1122
24.5043 | -103.1865
24.508 | -103.2624
24.5117 | -103.3397
24.5154 | -103.4185
24.519 | -103.4984
24.5227 | -103.5793
24.5264 | -103.6601
24.5301 | -103.7405
24.5337 | -103.8208
24.5374 | -103.9013
24.5411 | -103.98
24.5447 | -104.0567
24.5484 | -104.1306
24.5521 | -104.2025
24.5558 | -104.2731
24.5594 | -104.3401
24.5631 | -104.4001
24.5668 | -104.4554
24.5705 | -104.5018
24.5741 | -104.5357
24.5778 | -104.5667
24.5815 | -104.5963
24.5852 | -104.6276
24.5888 | -104.6589
24.5925 | -104.688
24.5962 | -104.7147
24.5998 | -104.7372
24.6035 | -104.7549
24.6072 | -104.7787
24.6109 | -104.8095
24.6145 | -104.8373
24.6182 | -104.8505
24.6219 | -104.8401
24.6256 | -104.8079
24.6292 | -104.7703
24.6329 | -104.7317
24.6366 | -104.6938
24.6403 | -104.6548
24.6439 | -104.613
24.6476 | -104.5694
24.6513 | -104.5244
24.6549 | -104.4775
24.6586 | -104.4304
24.6623 | -104.3828
24.666 | -104.3342
24.6696 | -104.2833
24.6733 | -104.2335
24.677 | -104.185
24.6807 | -104.1371
24.6843 | -104.0898
24.688 | -104.0426
24.6917 | -103.9962
24.6953 | -103.9509
24.699 | -103.9076
24.7027 | -103.8665
24.7064 | -103.8276
24.71 | -103.7907
24.7137 | -103.7558
24.7174 | -103.7223
24.7211 | -103.6913
24.7247 | -103.663
24.7284 | -103.638
24.7321 | -103.616
24.7358 | -103.5974
24.7394 | -103.583
24.7431 | -103.5708
24.7468 | -103.5604
24.7504 | -103.5547
24.7541 | -103.5532
24.7578 | -103.5549
24.7615 | -103.5594
24.7651 | -103.5659
24.7688 | -103.5758
24.7725 | -103.5889
24.7762 | -103.605
24.7798 | -103.6247
24.7835 | -103.6488
24.7872 | -103.6757
24.7909 | -103.704
24.7945 | -103.7355
24.7982 | -103.7713
24.8019 | -103.8094
24.8055 | -103.8481
24.8092 | -103.8887
24.8129 | -103.9313
24.8166 | -103.9761
24.8202 | -104.024
24.8239 | -104.0718
24.8276 | -104.1196
24.8313 | -104.1688
24.8349 | -104.2186
24.8386 | -104.2695
24.8423 | -104.3207
24.846 | -104.3716
24.8496 | -104.4218
24.8533 | -104.4701
24.857 | -104.5182
24.8606 | -104.5646
24.8643 | -104.6097
24.868 | -104.6549
24.8717 | -104.6978
24.8753 | -104.74
24.879 | -104.7803
24.8827 | -104.8207
24.8864 | -104.8587
24.89 | -104.8949
24.8937 | -104.9291
24.8974 | -104.9634
24.9011 | -104.9969
24.9047 | -105.0073
24.9084 | -105.0331
24.9121 | -105.0608
24.9157 | -105.0914
24.9194 | -105.1191
24.9231 | -105.1497
24.9268 | -105.1833
24.9304 | -105.2101
24.9341 | -105.2297
24.9378 | -105.2415
24.9415 | -105.2475
24.9451 | -105.2518
24.9488 | -105.2556
24.9525 | -105.2585
24.9562 | -105.2607
24.9598 | -105.2625
24.9635 | -105.2646
24.9672 | -105.2686
24.9708 | -105.2724
24.9745 | -105.2736
24.9782 | -105.2735
24.9819 | -105.2715
24.9855 | -105.2683
24.9892 | -105.2644
24.9929 | -105.259
24.9966 | -105.2543
25.0002 | -105.2496
25.0039 | -105.2449
25.0076 | -105.2408
25.0112 | -105.2358
25.0149 | -105.2323
25.0186 | -105.2291
25.0223 | -105.2264
25.0259 | -105.2226
25.0296 | -105.2183
25.0333 | -105.2131
25.037 | -105.2084
25.0406 | -105.2039
25.0443 | -105.2003
25.048 | -105.1973
25.0517 | -105.193
25.0553 | -105.1896
25.059 | -105.1851
25.0627 | -105.1803
25.0663 | -105.175
25.07 | -105.1716
25.0737 | -105.1696
25.0774 | -105.1656
25.081 | -105.1639
25.0847 | -105.1622
25.0884 | -105.1606
25.0921 | -105.1591
25.0957 | -105.1575
25.0994 | -105.1561
25.1031 | -105.1535
25.1068 | -105.1524
25.1104 | -105.1513
25.1141 | -105.1515
25.1178 | -105.1525
25.1214 | -105.1538
25.1251 | -105.1551
25.1288 | -105.1568
25.1325 | -105.1573
25.1361 | -105.1574
25.1398 | -105.1572
25.1435 | -105.1589
25.1472 | -105.1606
25.1508 | -105.1602
25.1545 | -105.1608
25.1582 | -105.1629
25.1619 | -105.1642
25.1655 | -105.1648
25.1692 | -105.1647
25.1729 | -105.1632
25.1765 | -105.1641
25.1802 | -105.1641
25.1839 | -105.1635
25.1876 | -105.1632
25.1912 | -105.1624
25.1949 | -105.1609
25.1986 | -105.1594
25.2023 | -105.1586
25.2059 | -105.1565
25.2096 | -105.1558
25.2133 | -105.1566
25.217 | -105.1574
25.2206 | -105.1555
25.2243 | -105.154
25.228 | -105.153
25.2316 | -105.1507
25.2353 | -105.1456
25.239 | -105.1414
25.2427 | -105.139
25.2463 | -105.137
25.25 | -105.1334
25.2537 | -105.1291
25.2574 | -105.1269
25.261 | -105.1252
25.2647 | -105.1245
25.2684 | -105.1218
25.2721 | -105.1184
25.2757 | -105.1151
25.2794 | -105.1122
25.2831 | -105.1099
25.2867 | -105.1065
25.2904 | -105.1035
25.2941 | -105.0986
25.2978 | -105.094
25.3014 | -105.0913
25.3051 | -105.0881
25.3088 | -105.0853
25.3125 | -105.0828
25.3161 | -105.0782
25.3198 | -105.073
25.3235 | -105.0683
25.3271 | -105.0633
25.3308 | -105.0591
25.3345 | -105.0557
25.3382 | -105.0524
25.3418 | -105.0495
25.3455 | -105.0463
25.3492 | -105.0438
25.3529 | -105.0426
25.3565 | -105.041
25.3602 | -105.0394
25.3639 | -105.0385
25.3676 | -105.0367
25.3712 | -105.035
25.3749 | -105.0347
25.3786 | -105.033
25.3822 | -105.033
25.3859 | -105.0322
25.3896 | -105.0329
25.3933 | -105.0332
25.3969 | -105.0325
25.4006 | -105.0315
25.4043 | -105.0322
25.408 | -105.0333
25.4116 | -105.0333
25.4153 | -105.0347
25.419 | -105.0376
25.4227 | -105.0405
25.4263 | -105.0407
25.43 | -105.0412
25.4337 | -105.0428
25.4373 | -105.0448
25.441 | -105.045
25.4447 | -105.0454
25.4484 | -105.0448
25.452 | -105.0447
25.4557 | -105.0458
25.4594 | -105.0457
25.4631 | -105.0459
25.4667 | -105.0452
25.4704 | -105.0442
25.4741 | -105.0422
25.4778 | -105.0388
25.4814 | -105.0345
25.4851 | -105.0287
25.4888 | -105.023
25.4924 | -105.0182
25.4961 | -105.0145
25.4998 | -105.0093
25.5035 | -105.0031
25.5071 | -104.9957
25.5108 | -104.988
25.5145 | -104.9807
25.5182 | -104.9731
25.5218 | -104.9657
25.5255 | -104.9581
25.5292 | -104.9508
25.5329 | -104.9439
25.5365 | -104.9354
25.5402 | -104.9277
25.5439 | -104.9193
25.5475 | -104.9105
25.5512 | -104.9021
25.5549 | -104.8937
25.5586 | -104.8858
25.5622 | -104.8783
25.5659 | -104.8721
25.5696 | -104.8644
25.5733 | -104.8568
25.5769 | -104.8504
25.5806 | -104.8452
25.5843 | -104.8388
25.588 | -104.8318
25.5916 | -104.828
25.5953 | -104.8261
25.599 | -104.8239
25.6026 | -104.82
25.6063 | -104.816
25.61 | -104.8133
25.6137 | -104.8126
25.6173 | -104.8116
25.621 | -104.8099
25.6247 | -104.8085
25.6284 | -104.8077
25.632 | -104.8072
25.6357 | -104.8075
25.6394 | -104.8069
25.6431 | -104.8073
25.6467 | -104.8089
25.6504 | -104.812
25.6541 | -104.8154
25.6577 | -104.8177
25.6614 | -104.8191
25.6651 | -104.8206
25.6688 | -104.8218
25.6724 | -104.8232
25.6761 | -104.8255
25.6798 | -104.8298
25.6835 | -104.8337
25.6871 | -104.8374
25.6908 | -104.8419
25.6945 | -104.8444
25.6981 | -104.8468
25.7018 | -104.8482
25.7055 | -104.8502
25.7092 | -104.8522
25.7128 | -104.8541
25.7165 | -104.8549
25.7202 | -104.8561
25.7239 | -104.8574
25.7275 | -104.8581
25.7312 | -104.8585
25.7349 | -104.8587
25.7386 | -104.8585
25.7422 | -104.8558
25.7459 | -104.8546
25.7496 | -104.8543
25.7532 | -104.8544
25.7569 | -104.8557
25.7606 | -104.8563
25.7643 | -104.8528
25.7679 | -104.8495
25.7716 | -104.8485
25.7753 | -104.8479
25.779 | -104.8478
25.7826 | -104.8468
25.7863 | -104.8476
25.79 | -104.8467
25.7937 | -104.8452
25.7973 | -104.8448
25.801 | -104.8457
25.8047 | -104.8463
25.8083 | -104.846
25.812 | -104.8459
25.8157 | -104.8466
25.8194 | -104.8471
25.823 | -104.8499
25.8267 | -104.8522
25.8304 | -104.8535
25.8341 | -104.8566
25.8377 | -104.8597
25.8414 | -104.8635
25.8451 | -104.868
25.8488 | -104.8745
25.8524 | -104.8809
25.8561 | -104.887
25.8598 | -104.8928
25.8634 | -104.8986
25.8671 | -104.9053
25.8708 | -104.9115
25.8745 | -104.9183
25.8781 | -104.9255
25.8818 | -104.9312
25.8855 | -104.9359
25.8892 | -104.9417
25.8928 | -104.9491
25.8965 | -104.957
25.9002 | -104.9654
25.9039 | -104.9739
25.9075 | -104.9824
25.9112 | -104.9908
25.9149 | -104.9991
25.9185 | -105.0084
25.9222 | -105.0169
25.9259 | -105.0247
25.9296 | -105.0324
25.9332 | -105.0407
25.9369 | -105.0489
25.9406 | -105.0565
25.9443 | -105.0645
25.9479 | -105.0744
25.9516 | -105.0835
25.9553 | -105.0917
25.959 | -105.0991
25.9626 | -105.1057
25.9663 | -105.1134
25.97 | -105.121
25.9736 | -105.1292
25.9773 | -105.1376
25.981 | -105.145
25.9847 | -105.1523
25.9883 | -105.1614
25.992 | -105.1706
25.9957 | -105.1792
25.9994 | -105.188
26.003 | -105.1968
26.0067 | -105.2063
26.0104 | -105.2146
26.0141 | -105.2232
26.0177 | -105.2317
26.0214 | -105.2406
26.0251 | -105.2472
26.0287 | -105.2555
26.0324 | -105.2657
26.0361 | -105.275
26.0398 | -105.2833
26.0434 | -105.2904
26.0471 | -105.2964
26.0508 | -105.3027
26.0545 | -105.31
26.0581 | -105.3164
26.0618 | -105.3246
26.0655 | -105.333
26.0691 | -105.3411
26.0728 | -105.349
26.0765 | -105.3558
26.0802 | -105.3627
26.0838 | -105.3714
26.0875 | -105.3802
26.0912 | -105.3872
26.0949 | -105.3956
26.0985 | -105.4034
26.1022 | -105.4109
26.1059 | -105.4183
26.1096 | -105.4257
26.1132 | -105.4321
26.1169 | -105.4365
26.1206 | -105.4408
26.1242 | -105.4426
26.1279 | -105.4459
26.1316 | -105.453
26.1353 | -105.4587
26.1389 | -105.4607
26.1426 | -105.4619
26.1463 | -105.4615
26.15 | -105.4626
26.1536 | -105.4641
26.1573 | -105.4655
26.161 | -105.4646
26.1647 | -105.4653
26.1683 | -105.4659
26.172 | -105.4664
26.1757 | -105.4668
26.1793 | -105.4678
26.183 | -105.4678
26.1867 | -105.4672
26.1904 | -105.465
26.194 | -105.4625
26.1977 | -105.4603
26.2014 | -105.4581
26.2051 | -105.4552
26.2087 | -105.4521
26.2124 | -105.4489
26.2161 | -105.4463
26.2198 | -105.444
26.2234 | -105.4418
26.2271 | -105.4381
26.2308 | -105.4353
26.2344 | -105.4312
26.2381 | -105.4272
26.2418 | -105.423
26.2455 | -105.4193
26.2491 | -105.4184
26.2528 | -105.4172
26.2565 | -105.4151
26.2602 | -105.4105
26.2638 | -105.406
26.2675 | -105.4019
26.2712 | -105.3991
26.2749 | -105.3954
26.2785 | -105.3913
26.2822 | -105.3876
26.2859 | -105.3832
26.2895 | -105.3797
26.2932 | -105.376
26.2969 | -105.3711
26.3006 | -105.367
26.3042 | -105.3633
26.3079 | -105.3591
26.3116 | -105.3534
26.3153 | -105.3473
26.3189 | -105.3429
26.3226 | -105.338
26.3263 | -105.3336
26.33 | -105.3295
26.3336 | -105.3257
26.3373 | -105.321
26.341 | -105.3164
26.3446 | -105.3139
26.3483 | -105.3103
26.352 | -105.3054
26.3557 | -105.3008
26.3593 | -105.2966
26.363 | -105.293
26.3667 | -105.2904
26.3704 | -105.2873
26.374 | -105.284
26.3777 | -105.282
26.3814 | -105.2801
26.385 | -105.279
26.3887 | -105.2764
26.3924 | -105.2731
26.3961 | -105.2722
26.3997 | -105.2703
26.4034 | -105.2693
26.4071 | -105.2699
26.4108 | -105.2696
26.4144 | -105.272
26.4181 | -105.2756
26.4218 | -105.2786
26.4255 | -105.2806
26.4291 | -105.2832
26.4328 | -105.2864
26.4365 | -105.2895
26.4401 | -105.293
26.4438 | -105.2973
26.4475 | -105.3036
26.4512 | -105.3104
26.4548 | -105.316
26.4585 | -105.3217
26.4622 | -105.3293
26.4659 | -105.3346
26.4695 | -105.3394
26.4732 | -105.3445
26.4769 | -105.3518
26.4806 | -105.3598
26.4842 | -105.3661
26.4879 | -105.3722
26.4916 | -105.3787
26.4952 | -105.3857
26.4989 | -105.3907
26.5026 | -105.3954
26.5063 | -105.4004
26.5099 | -105.4051
26.5136 | -105.4099
26.5173 | -105.4158
26.521 | -105.4208
26.5246 | -105.426
26.5283 | -105.4284
26.532 | -105.4305
26.5357 | -105.4341
26.5393 | -105.4348
26.543 | -105.4351
26.5467 | -105.4356
26.5503 | -105.436
26.554 | -105.437
26.5577 | -105.4375
26.5614 | -105.4377
26.565 | -105.4367
26.5687 | -105.4358
26.5724 | -105.4359
26.5761 | -105.4352
26.5797 | -105.4339
26.5834 | -105.433
26.5871 | -105.432
26.5908 | -105.4289
26.5944 | -105.4264
26.5981 | -105.4243
26.6018 | -105.4239
26.6054 | -105.423
26.6091 | -105.4206
26.6128 | -105.4197
26.6165 | -105.4178
26.6201 | -105.414
26.6238 | -105.4094
26.6275 | -105.4054
26.6312 | -105.4017
26.6348 | -105.4004
26.6385 | -105.3977
26.6422 | -105.3932
26.6459 | -105.3885
26.6495 | -105.3865
26.6532 | -105.3839
26.6569 | -105.3791
26.6605 | -105.3756
26.6642 | -105.3722
26.6679 | -105.3691
26.6716 | -105.367
26.6752 | -105.365
26.6789 | -105.3617
26.6826 | -105.3596
26.6863 | -105.3598
26.6899 | -105.3599
26.6936 | -105.3578
26.6973 | -105.3559
26.7009 | -105.3537
26.7046 | -105.3519
26.7083 | -105.3498
26.712 | -105.3495
26.7156 | -105.3493
26.7193 | -105.3466
26.723 | -105.3443
26.7267 | -105.3435
26.7303 | -105.3427
26.734 | -105.3427
26.7377 | -105.3425
26.7414 | -105.3424
26.745 | -105.3408
26.7487 | -105.3402
26.7524 | -105.3405
26.756 | -105.3404
26.7597 | -105.34
26.7634 | -105.3396
26.7671 | -105.3412
26.7707 | -105.3438
26.7744 | -105.3462
26.7781 | -105.3468
26.7818 | -105.3476
26.7854 | -105.3496
26.7891 | -105.3492
26.7928 | -105.3488
26.7965 | -105.3489
26.8001 | -105.3484
26.8038 | -105.3477
26.8075 | -105.3471
26.8111 | -105.3476
26.8148 | -105.3485
26.8185 | -105.3492
26.8222 | -105.3492
26.8258 | -105.3476
26.8295 | -105.3471
26.8332 | -105.3469
26.8369 | -105.3462
26.8405 | -105.3472
26.8442 | -105.3499
26.8479 | -105.3524
26.8516 | -105.3532
26.8552 | -105.3542
26.8589 | -105.3549
26.8626 | -105.3552
26.8662 | -105.3565
26.8699 | -105.3569
26.8736 | -105.3581
26.8773 | -105.3621
26.8809 | -105.3648
26.8846 | -105.3659
26.8883 | -105.3675
26.892 | -105.3703
26.8956 | -105.3726
26.8993 | -105.3756
26.903 | -105.3781
26.9067 | -105.3814
26.9103 | -105.3854
26.914 | -105.3904
26.9177 | -105.3949
26.9213 | -105.3998
26.925 | -105.4054
26.9287 | -105.4106
26.9324 | -105.4134
26.936 | -105.4174
26.9397 | -105.4224
26.9434 | -105.428
26.9471 | -105.4343
26.9507 | -105.4383
26.9544 | -105.4424
26.9581 | -105.4481
26.9618 | -105.4528
26.9654 | -105.4581
26.9691 | -105.4619
26.9728 | -105.4658
26.9764 | -105.4692
26.9801 | -105.4727
26.9838 | -105.4767
26.9875 | -105.4827
26.9911 | -105.4868
26.9948 | -105.4897
26.9985 | -105.4936
27.0022 | -105.499
27.0058 | -105.5042
27.0095 | -105.5079
27.0132 | -105.509
27.0169 | -105.5099
27.0205 | -105.5121
27.0242 | -105.5123
27.0279 | -105.5145
27.0315 | -105.515
27.0352 | -105.5167
27.0389 | -105.5201
27.0426 | -105.5224
27.0462 | -105.5238
27.0499 | -105.5247
27.0536 | -105.5252
27.0573 | -105.5245
27.0609 | -105.5252
27.0646 | -105.525
27.0683 | -105.5241
27.072 | -105.5235
27.0756 | -105.5234
27.0793 | -105.5242
27.083 | -105.5249
27.0866 | -105.525
27.0903 | -105.5242
27.094 | -105.5225
27.0977 | -105.522
27.1013 | -105.5237
27.105 | -105.5258
27.1087 | -105.5265
27.1124 | -105.5256
27.116 | -105.5245
27.1197 | -105.5249
27.1234 | -105.5247
27.127 | -105.5274
27.1307 | -105.5302
27.1344 | -105.53
27.1381 | -105.5305
27.1417 | -105.5326
27.1454 | -105.5353
27.1491 | -105.536
27.1528 | -105.5369
27.1564 | -105.537
27.1601 | -105.5387
27.1638 | -105.5419
27.1675 | -105.5453
27.1711 | -105.5502
27.1748 | -105.5528
27.1785 | -105.5538
27.1821 | -105.5568
27.1858 | -105.5606
27.1895 | -105.5642
27.1932 | -105.5659
27.1968 | -105.5676
27.2005 | -105.57
27.2042 | -105.5723
27.2079 | -105.5747
27.2115 | -105.5775
27.2152 | -105.5792
27.2189 | -105.5803
27.2226 | -105.5821
27.2262 | -105.5839
27.2299 | -105.5835
27.2336 | -105.5825
27.2372 | -105.5818
27.2409 | -105.5835
27.2446 | -105.5861
27.2483 | -105.5862
27.2519 | -105.5846
27.2556 | -105.5837
27.2593 | -105.5832
27.263 | -105.582
27.2666 | -105.5801
27.2703 | -105.578
27.274 | -105.5745
27.2777 | -105.5703
27.2813 | -105.5666
27.285 | -105.5627
27.2887 | -105.5597
27.2923 | -105.5568
27.296 | -105.5512
27.2997 | -105.5461
27.3034 | -105.5404
27.307 | -105.5354
27.3107 | -105.5297
27.3144 | -105.5235
27.3181 | -105.517
27.3217 | -105.5106
27.3254 | -105.505
27.3291 | -105.498
27.3328 | -105.4906
27.3364 | -105.4828
27.3401 | -105.4768
27.3438 | -105.4707
27.3474 | -105.4644
27.3511 | -105.4606
27.3548 | -105.4544
27.3585 | -105.4483
27.3621 | -105.4415
27.3658 | -105.436
27.3695 | -105.4305
27.3732 | -105.4245
27.3768 | -105.419
27.3805 | -105.4139
27.3842 | -105.4074
27.3879 | -105.4017
27.3915 | -105.3975
27.3952 | -105.3952
27.3989 | -105.3941
27.4025 | -105.3922
27.4062 | -105.3889
27.4099 | -105.3859
27.4136 | -105.3849
27.4172 | -105.3833
27.4209 | -105.3827
27.4246 | -105.3827
27.4283 | -105.382
27.4319 | -105.3801
27.4356 | -105.3806
27.4393 | -105.3819
27.4429 | -105.3835
27.4466 | -105.3843
27.4503 | -105.3855
27.454 | -105.3867
27.4576 | -105.3868
27.4613 | -105.3877
27.465 | -105.387
27.4687 | -105.3864
27.4723 | -105.3869
27.476 | -105.3884
27.4797 | -105.3891
27.4834 | -105.3903
27.487 | -105.3911
27.4907 | -105.3918
27.4944 | -105.3933
27.498 | -105.396
27.5017 | -105.3972
27.5054 | -105.3963
27.5091 | -105.3966
27.5127 | -105.3962
27.5164 | -105.3949
27.5201 | -105.3959
27.5238 | -105.3966
27.5274 | -105.3949
27.5311 | -105.392
27.5348 | -105.3904
27.5385 | -105.389
27.5421 | -105.3876
27.5458 | -105.3862
27.5495 | -105.3827
27.5531 | -105.3801
27.5568 | -105.3773
27.5605 | -105.3739
27.5642 | -105.373
27.5678 | -105.3739
27.5715 | -105.3723
27.5752 | -105.3698
27.5789 | -105.3672
27.5825 | -105.3666
27.5862 | -105.3652
27.5899 | -105.3634
27.5936 | -105.3617
27.5972 | -105.3607
27.6009 | -105.3581
27.6046 | -105.3561
27.6082 | -105.3557
27.6119 | -105.3548
27.6156 | -105.3534
27.6193 | -105.3545
27.6229 | -105.3559
27.6266 | -105.3575
27.6303 | -105.3585
27.634 | -105.3585
27.6376 | -105.3592
27.6413 | -105.3589
27.645 | -105.3588
27.6487 | -105.3592
27.6523 | -105.3608
27.656 | -105.3637
27.6597 | -105.3663
27.6633 | -105.3694
27.667 | -105.3733
27.6707 | -105.3775
27.6744 | -105.3818
27.678 | -105.3853
27.6817 | -105.3883
27.6854 | -105.3927
27.6891 | -105.3976
27.6927 | -105.402
27.6964 | -105.4049
27.7001 | -105.4089
27.7038 | -105.4142
27.7074 | -105.4183
27.7111 | -105.422
27.7148 | -105.4262
27.7184 | -105.4302
27.7221 | -105.4372
27.7258 | -105.4441
27.7295 | -105.4483
27.7331 | -105.4524
27.7368 | -105.4565
27.7405 | -105.4593
27.7442 | -105.4637
27.7478 | -105.4684
27.7515 | -105.4728
27.7552 | -105.4757
27.7588 | -105.4772
27.7625 | -105.4797
27.7662 | -105.4839
27.7699 | -105.4864
27.7735 | -105.4886
27.7772 | -105.4905
27.7809 | -105.4916
27.7846 | -105.4898
27.7882 | -105.4874
27.7919 | -105.4865
27.7956 | -105.4863
27.7993 | -105.4855
27.8029 | -105.4832
27.8066 | -105.481
27.8103 | -105.4776
27.8139 | -105.4743
27.8176 | -105.4727
27.8213 | -105.4707
27.825 | -105.4666
27.8286 | -105.4632
27.8323 | -105.4594
27.836 | -105.4541
27.8397 | -105.4489
27.8433 | -105.4447
27.847 | -105.4411
27.8507 | -105.4352
27.8544 | -105.4293
27.858 | -105.4235
27.8617 | -105.4164
27.8654 | -105.4102
27.869 | -105.4044
27.8727 | -105.3977
27.8764 | -105.3898
27.8801 | -105.3824
27.8837 | -105.3758
27.8874 | -105.3678
27.8911 | -105.3611
27.8948 | -105.3557
27.8984 | -105.3493
27.9021 | -105.3438
27.9058 | -105.3391
27.9095 | -105.3335
27.9131 | -105.3284
27.9168 | -105.324
27.9205 | -105.3206
27.9241 | -105.3145
27.9278 | -105.3066
27.9315 | -105.2998
27.9352 | -105.2933
27.9388 | -105.2893
27.9425 | -105.2862
27.9462 | -105.2816
27.9499 | -105.2766
27.9535 | -105.2738
27.9572 | -105.2711
27.9609 | -105.2685
27.9646 | -105.2678
27.9682 | -105.2673
27.9719 | -105.2664
27.9756 | -105.2648
27.9792 | -105.264
27.9829 | -105.2625
27.9866 | -105.2613
27.9903 | -105.2623
27.9939 | -105.2641
27.9976 | -105.2655
28.0013 | -105.2657