## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 283 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 286
Hood | 02-92B/P | 220
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1012 | hPa |  | PRC PRO 01
Humidity | 55.2 | %/Hr
Internal temperature of calibration bench | 24.01 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1018 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.5 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.92 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.01 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3067 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 03/09/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 05/09/2018
Technical Manager | Anthony HUE | Approval |  | 27/09/2018
CEO | Robert GLEMOT | Approval |  | 27/09/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4565 | 57.4384
0.012209 | -36.0325 | 38.2829
0.018313 | -34.9799 | 28.1561
0.024417 | -34.5138 | 21.9262
0.030521 | -34.4033 | 18.1431
0.036626 | -34.2591 | 15.0996
0.04273 | -34.1642 | 12.7598
0.048834 | -34.119 | 11.241
0.054938 | -34.0549 | 10.1046
0.061043 | -33.9951 | 9.0306
0.067147 | -33.9855 | 8.2082
0.073251 | -33.9868 | 7.4438
0.079355 | -34.0045 | 6.7757
0.08546 | -34.0174 | 6.1795
0.091564 | -34.012 | 5.7396
0.097668 | -33.9944 | 5.4909
0.10377 | -33.9734 | 5.2498
0.10988 | -33.9565 | 4.98
0.11598 | -33.9469 | 4.695
0.12209 | -33.9477 | 4.4147
0.12819 | -33.9567 | 4.1291
0.13429 | -33.9672 | 3.8346
0.1404 | -33.9706 | 3.5538
0.1465 | -33.9728 | 3.2957
0.15261 | -33.9707 | 3.0783
0.15871 | -33.9653 | 2.9216
0.16482 | -33.958 | 2.8258
0.17092 | -33.955 | 2.7901
0.17702 | -33.9523 | 2.7203
0.18313 | -33.9506 | 2.6422
0.18923 | -33.9496 | 2.5322
0.19534 | -33.9481 | 2.4042
0.20144 | -33.9471 | 2.2615
0.20754 | -33.9452 | 2.1433
0.21365 | -33.9448 | 2.0315
0.21975 | -33.9448 | 1.9367
0.22586 | -33.946 | 1.8532
0.23196 | -33.9484 | 1.7835
0.23807 | -33.9521 | 1.7221
0.24417 | -33.9559 | 1.6607
0.25027 | -33.9596 | 1.5963
0.25638 | -33.962 | 1.5322
0.26248 | -33.9639 | 1.4634
0.26859 | -33.9649 | 1.3967
0.27469 | -33.9653 | 1.3256
0.2808 | -33.9663 | 1.2539
0.2869 | -33.9669 | 1.1837
0.293 | -33.9676 | 1.1172
0.29911 | -33.9684 | 1.0574
0.30521 | -33.9689 | 0.99833
0.31132 | -33.9692 | 0.95881
0.31742 | -33.9701 | 0.91576
0.32353 | -33.97 | 0.87964
0.32963 | -33.9704 | 0.84964
0.33573 | -33.9711 | 0.8284
0.34184 | -33.9722 | 0.8066
0.34794 | -33.9742 | 0.78266
0.35405 | -33.9761 | 0.75821
0.36015 | -33.9786 | 0.73208
0.36626 | -33.9806 | 0.70446
0.37236 | -33.9828 | 0.66795
0.37846 | -33.9852 | 0.62414
0.38457 | -33.9863 | 0.56864
0.39067 | -33.9875 | 0.52369
0.39678 | -33.9884 | 0.47386
0.40288 | -33.9892 | 0.4222
0.40899 | -33.9894 | 0.37234
0.41509 | -33.9891 | 0.32526
0.42119 | -33.9895 | 0.28608
0.4273 | -33.9898 | 0.24306
0.4334 | -33.9901 | 0.20067
0.43951 | -33.9906 | 0.1607
0.44561 | -33.9918 | 0.12349
0.45172 | -33.9934 | 0.092807
0.45782 | -33.9943 | 0.056577
0.46392 | -33.9957 | 0.021293
0.47003 | -33.9967 | -0.011519
0.47613 | -33.9976 | -0.046346
0.48224 | -33.9981 | -0.0823
0.48834 | -33.9971 | -0.11993
0.49445 | -33.9969 | -0.15737
0.50055 | -33.9962 | -0.1916
0.50665 | -33.9956 | -0.22883
0.51276 | -33.9946 | -0.26437
0.51886 | -33.9936 | -0.30043
0.52497 | -33.9923 | -0.32537
0.53107 | -33.9915 | -0.3562
0.53717 | -33.9908 | -0.38543
0.54328 | -33.9898 | -0.40895
0.54938 | -33.9895 | -0.43254
0.55549 | -33.9894 | -0.45415
0.56159 | -33.9902 | -0.46635
0.5677 | -33.9902 | -0.48184
0.5738 | -33.9909 | -0.49866
0.5799 | -33.9913 | -0.51409
0.58601 | -33.9921 | -0.52498
0.59211 | -33.9934 | -0.53264
0.59822 | -33.9945 | -0.54251
0.60432 | -33.9958 | -0.55537
0.61043 | -33.9969 | -0.56401
0.61653 | -33.9979 | -0.57462
0.62263 | -33.9993 | -0.58362
0.62874 | -34.0001 | -0.5926
0.63484 | -34.0009 | -0.60132
0.64095 | -34.0014 | -0.61011
0.64705 | -34.0018 | -0.61812
0.65316 | -34.0021 | -0.62896
0.65926 | -34.0024 | -0.64298
0.66536 | -34.0015 | -0.65875
0.67147 | -34.0014 | -0.67012
0.67757 | -34.0012 | -0.68289
0.68368 | -34.0007 | -0.6998
0.68978 | -34.0002 | -0.71507
0.69589 | -33.9994 | -0.7332
0.70199 | -33.9986 | -0.75576
0.70809 | -33.9978 | -0.77438
0.7142 | -33.9973 | -0.79379
0.7203 | -33.9963 | -0.81271
0.72641 | -33.9957 | -0.83517
0.73251 | -33.995 | -0.85499
0.73862 | -33.9951 | -0.87834
0.74472 | -33.9947 | -0.8992
0.75082 | -33.9943 | -0.92201
0.75693 | -33.9946 | -0.94465
0.76303 | -33.995 | -0.96697
0.76914 | -33.9966 | -0.9884
0.77524 | -33.9974 | -1.0086
0.78135 | -33.9985 | -1.0294
0.78745 | -33.9996 | -1.0501
0.79355 | -34.0006 | -1.0704
0.79966 | -34.0021 | -1.0906
0.80576 | -34.0043 | -1.1102
0.81187 | -34.0057 | -1.1324
0.81797 | -34.0075 | -1.1541
0.82408 | -34.0091 | -1.1753
0.83018 | -34.0107 | -1.1962
0.83628 | -34.0125 | -1.2149
0.84239 | -34.0137 | -1.235
0.84849 | -34.015 | -1.2569
0.8546 | -34.0164 | -1.2763
0.8607 | -34.0176 | -1.2967
0.86681 | -34.0186 | -1.3154
0.87291 | -34.0196 | -1.3337
0.87901 | -34.0201 | -1.3514
0.88512 | -34.0209 | -1.3692
0.89122 | -34.0217 | -1.3871
0.89733 | -34.0221 | -1.4025
0.90343 | -34.0226 | -1.4201
0.90953 | -34.0228 | -1.4341
0.91564 | -34.0229 | -1.4498
0.92174 | -34.0231 | -1.4646
0.92785 | -34.0233 | -1.4791
0.93395 | -34.0235 | -1.493
0.94006 | -34.0237 | -1.5059
0.94616 | -34.0234 | -1.5182
0.95226 | -34.0236 | -1.5309
0.95837 | -34.0237 | -1.5447
0.96447 | -34.0239 | -1.5573
0.97058 | -34.0238 | -1.5708
0.97668 | -34.0236 | -1.5843
0.98279 | -34.0234 | -1.597
0.98889 | -34.0232 | -1.6079
0.99499 | -34.0232 | -1.6218
1.0011 | -34.0232 | -1.6323
1.0072 | -34.0229 | -1.6441
1.0133 | -34.0226 | -1.655
1.0194 | -34.0222 | -1.668
1.0255 | -34.022 | -1.6817
1.0316 | -34.0216 | -1.6938
1.0377 | -34.0212 | -1.7073
1.0438 | -34.0207 | -1.72
1.0499 | -34.02 | -1.7353
1.056 | -34.0195 | -1.7496
1.0621 | -34.019 | -1.7635
1.0682 | -34.0186 | -1.7758
1.0743 | -34.018 | -1.791
1.0805 | -34.0176 | -1.8051
1.0866 | -34.0167 | -1.8195
1.0927 | -34.0158 | -1.8338
1.0988 | -34.0152 | -1.8486
1.1049 | -34.0144 | -1.8626
1.111 | -34.0137 | -1.8778
1.1171 | -34.013 | -1.8921
1.1232 | -34.0125 | -1.9069
1.1293 | -34.0121 | -1.9209
1.1354 | -34.0118 | -1.9357
1.1415 | -34.0114 | -1.951
1.1476 | -34.0107 | -1.9661
1.1537 | -34.0101 | -1.9807
1.1598 | -34.0101 | -1.9954
1.1659 | -34.0098 | -2.0087
1.172 | -34.0095 | -2.0224
1.1781 | -34.0092 | -2.0361
1.1842 | -34.009 | -2.05
1.1903 | -34.0092 | -2.0611
1.1964 | -34.0091 | -2.0747
1.2025 | -34.0091 | -2.0886
1.2086 | -34.0094 | -2.1011
1.2147 | -34.0093 | -2.1136
1.2209 | -34.0096 | -2.1249
1.227 | -34.0102 | -2.1359
1.2331 | -34.0106 | -2.1482
1.2392 | -34.0109 | -2.1603
1.2453 | -34.0112 | -2.1715
1.2514 | -34.0116 | -2.1819
1.2575 | -34.0121 | -2.1929
1.2636 | -34.0129 | -2.2022
1.2697 | -34.0135 | -2.2117
1.2758 | -34.0141 | -2.2229
1.2819 | -34.0148 | -2.2329
1.288 | -34.0154 | -2.2439
1.2941 | -34.0162 | -2.2523
1.3002 | -34.0168 | -2.2604
1.3063 | -34.0173 | -2.2697
1.3124 | -34.0179 | -2.279
1.3185 | -34.0187 | -2.288
1.3246 | -34.0193 | -2.2972
1.3307 | -34.0201 | -2.306
1.3368 | -34.0209 | -2.3149
1.3429 | -34.0215 | -2.3238
1.349 | -34.022 | -2.3331
1.3551 | -34.0227 | -2.3427
1.3613 | -34.0232 | -2.3511
1.3674 | -34.0238 | -2.3612
1.3735 | -34.0242 | -2.3709
1.3796 | -34.0246 | -2.3798
1.3857 | -34.0249 | -2.3892
1.3918 | -34.0254 | -2.3995
1.3979 | -34.026 | -2.4088
1.404 | -34.0264 | -2.42
1.4101 | -34.0268 | -2.4305
1.4162 | -34.0272 | -2.4412
1.4223 | -34.0276 | -2.4527
1.4284 | -34.0281 | -2.4628
1.4345 | -34.0283 | -2.4743
1.4406 | -34.0286 | -2.4879
1.4467 | -34.029 | -2.4994
1.4528 | -34.0294 | -2.5106
1.4589 | -34.0299 | -2.5223
1.465 | -34.0301 | -2.5332
1.4711 | -34.0305 | -2.5446
1.4772 | -34.031 | -2.5581
1.4833 | -34.0314 | -2.5711
1.4894 | -34.0319 | -2.5834
1.4955 | -34.0325 | -2.5951
1.5016 | -34.033 | -2.6072
1.5078 | -34.0334 | -2.6211
1.5139 | -34.0338 | -2.6338
1.52 | -34.0342 | -2.6461
1.5261 | -34.0345 | -2.6587
1.5322 | -34.0349 | -2.6723
1.5383 | -34.0353 | -2.6846
1.5444 | -34.0358 | -2.6967
1.5505 | -34.0363 | -2.71
1.5566 | -34.0368 | -2.7228
1.5627 | -34.0373 | -2.7363
1.5688 | -34.038 | -2.7496
1.5749 | -34.0383 | -2.7618
1.581 | -34.0387 | -2.7746
1.5871 | -34.0391 | -2.7879
1.5932 | -34.0394 | -2.8004
1.5993 | -34.0399 | -2.8143
1.6054 | -34.0403 | -2.8274
1.6115 | -34.0407 | -2.8407
1.6176 | -34.041 | -2.8528
1.6237 | -34.0413 | -2.8659
1.6298 | -34.0416 | -2.878
1.6359 | -34.042 | -2.8905
1.642 | -34.0425 | -2.902
1.6482 | -34.0427 | -2.9149
1.6543 | -34.0431 | -2.9281
1.6604 | -34.0435 | -2.9398
1.6665 | -34.0436 | -2.9521
1.6726 | -34.0439 | -2.9644
1.6787 | -34.0442 | -2.9769
1.6848 | -34.0442 | -2.9893
1.6909 | -34.0444 | -3.0021
1.697 | -34.0447 | -3.0139
1.7031 | -34.0448 | -3.0266
1.7092 | -34.0448 | -3.0386
1.7153 | -34.045 | -3.0502
1.7214 | -34.0449 | -3.0622
1.7275 | -34.0448 | -3.0748
1.7336 | -34.0448 | -3.0876
1.7397 | -34.045 | -3.0999
1.7458 | -34.045 | -3.1118
1.7519 | -34.0451 | -3.1238
1.758 | -34.0452 | -3.1358
1.7641 | -34.0453 | -3.1474
1.7702 | -34.0453 | -3.1602
1.7763 | -34.0453 | -3.1717
1.7824 | -34.0453 | -3.1843
1.7885 | -34.0452 | -3.197
1.7947 | -34.0453 | -3.209
1.8008 | -34.0453 | -3.2204
1.8069 | -34.0453 | -3.2321
1.813 | -34.0454 | -3.2444
1.8191 | -34.0453 | -3.257
1.8252 | -34.0454 | -3.2699
1.8313 | -34.0455 | -3.2824
1.8374 | -34.0457 | -3.2944
1.8435 | -34.0457 | -3.3066
1.8496 | -34.0457 | -3.3199
1.8557 | -34.0459 | -3.3327
1.8618 | -34.0462 | -3.345
1.8679 | -34.0462 | -3.3571
1.874 | -34.0463 | -3.369
1.8801 | -34.0465 | -3.3817
1.8862 | -34.0467 | -3.3936
1.8923 | -34.0468 | -3.406
1.8984 | -34.0469 | -3.4182
1.9045 | -34.0469 | -3.43
1.9106 | -34.0469 | -3.442
1.9167 | -34.0469 | -3.4544
1.9228 | -34.047 | -3.4665
1.9289 | -34.0471 | -3.4784
1.9351 | -34.0472 | -3.4894
1.9412 | -34.0474 | -3.5008
1.9473 | -34.0476 | -3.5128
1.9534 | -34.0478 | -3.5243
1.9595 | -34.0479 | -3.5363
1.9656 | -34.0481 | -3.5484
1.9717 | -34.0483 | -3.5604
1.9778 | -34.0486 | -3.5719
1.9839 | -34.0488 | -3.5836
1.99 | -34.0489 | -3.5945
1.9961 | -34.0489 | -3.6069
2.0022 | -34.0492 | -3.6185
2.0083 | -34.0492 | -3.6299
2.0144 | -34.0493 | -3.6415
2.0205 | -34.0496 | -3.6534
2.0266 | -34.0499 | -3.6648
2.0327 | -34.0501 | -3.6764
2.0388 | -34.0505 | -3.6881
2.0449 | -34.0507 | -3.6996
2.051 | -34.0509 | -3.711
2.0571 | -34.0512 | -3.7224
2.0632 | -34.0514 | -3.7338
2.0693 | -34.0515 | -3.7461
2.0754 | -34.0518 | -3.7574
2.0816 | -34.0521 | -3.7697
2.0877 | -34.0522 | -3.7815
2.0938 | -34.0526 | -3.7936
2.0999 | -34.0527 | -3.8053
2.106 | -34.0529 | -3.8172
2.1121 | -34.0531 | -3.829
2.1182 | -34.0533 | -3.8397
2.1243 | -34.0536 | -3.8507
2.1304 | -34.0538 | -3.8617
2.1365 | -34.0541 | -3.8735
2.1426 | -34.0543 | -3.8854
2.1487 | -34.0546 | -3.8967
2.1548 | -34.055 | -3.9083
2.1609 | -34.0554 | -3.9197
2.167 | -34.0558 | -3.9309
2.1731 | -34.0561 | -3.9429
2.1792 | -34.0563 | -3.9537
2.1853 | -34.0567 | -3.9654
2.1914 | -34.057 | -3.9766
2.1975 | -34.0573 | -3.988
2.2036 | -34.0576 | -4
2.2097 | -34.0579 | -4.0111
2.2158 | -34.0582 | -4.0213
2.222 | -34.0585 | -4.0315
2.2281 | -34.0588 | -4.0426
2.2342 | -34.0591 | -4.0536
2.2403 | -34.0593 | -4.0638
2.2464 | -34.0596 | -4.0738
2.2525 | -34.0599 | -4.0848
2.2586 | -34.0602 | -4.0949
2.2647 | -34.0605 | -4.1052
2.2708 | -34.0608 | -4.116
2.2769 | -34.0612 | -4.126
2.283 | -34.0614 | -4.1365
2.2891 | -34.0616 | -4.1468
2.2952 | -34.0618 | -4.1569
2.3013 | -34.0621 | -4.1674
2.3074 | -34.0623 | -4.1772
2.3135 | -34.0626 | -4.188
2.3196 | -34.0628 | -4.1972
2.3257 | -34.0631 | -4.2072
2.3318 | -34.0632 | -4.2175
2.3379 | -34.0634 | -4.2276
2.344 | -34.0636 | -4.2382
2.3501 | -34.0638 | -4.2482
2.3562 | -34.064 | -4.2576
2.3623 | -34.0641 | -4.2671
2.3685 | -34.0642 | -4.2771
2.3746 | -34.0644 | -4.2875
2.3807 | -34.0647 | -4.2971
2.3868 | -34.0649 | -4.307
2.3929 | -34.0651 | -4.3167
2.399 | -34.0652 | -4.3264
2.4051 | -34.0654 | -4.3361
2.4112 | -34.0656 | -4.346
2.4173 | -34.0658 | -4.3558
2.4234 | -34.0659 | -4.3656
2.4295 | -34.0661 | -4.3757
2.4356 | -34.0663 | -4.3856
2.4417 | -34.0664 | -4.3959
2.4478 | -34.0666 | -4.4061
2.4539 | -34.0667 | -4.4158
2.46 | -34.0668 | -4.4265
2.4661 | -34.067 | -4.4369
2.4722 | -34.0672 | -4.4463
2.4783 | -34.0673 | -4.4561
2.4844 | -34.0674 | -4.4662
2.4905 | -34.0675 | -4.4765
2.4966 | -34.0677 | -4.4866
2.5027 | -34.0678 | -4.4965
2.5089 | -34.0679 | -4.5067
2.515 | -34.068 | -4.5168
2.5211 | -34.0681 | -4.5269
2.5272 | -34.0682 | -4.5373
2.5333 | -34.0684 | -4.5477
2.5394 | -34.0686 | -4.5578
2.5455 | -34.0687 | -4.5684
2.5516 | -34.0689 | -4.5782
2.5577 | -34.0691 | -4.5891
2.5638 | -34.0693 | -4.6006
2.5699 | -34.0693 | -4.6115
2.576 | -34.0694 | -4.6226
2.5821 | -34.0695 | -4.6332
2.5882 | -34.0696 | -4.6437
2.5943 | -34.0698 | -4.6543
2.6004 | -34.07 | -4.6651
2.6065 | -34.0701 | -4.6751
2.6126 | -34.0702 | -4.6852
2.6187 | -34.0704 | -4.6952
2.6248 | -34.0705 | -4.7057
2.6309 | -34.0707 | -4.7168
2.637 | -34.0709 | -4.7283
2.6431 | -34.0711 | -4.7385
2.6492 | -34.0712 | -4.7497
2.6554 | -34.0714 | -4.7605
2.6615 | -34.0715 | -4.7711
2.6676 | -34.0718 | -4.7823
2.6737 | -34.072 | -4.7936
2.6798 | -34.0723 | -4.8045
2.6859 | -34.0725 | -4.8155
2.692 | -34.0726 | -4.8263
2.6981 | -34.0728 | -4.8373
2.7042 | -34.0731 | -4.8486
2.7103 | -34.0733 | -4.8592
2.7164 | -34.0736 | -4.87
2.7225 | -34.0738 | -4.8813
2.7286 | -34.0741 | -4.8925
2.7347 | -34.0743 | -4.903
2.7408 | -34.0746 | -4.9133
2.7469 | -34.0748 | -4.9245
2.753 | -34.0751 | -4.9353
2.7591 | -34.0754 | -4.9466
2.7652 | -34.0757 | -4.9581
2.7713 | -34.0758 | -4.9689
2.7774 | -34.0761 | -4.9801
2.7835 | -34.0762 | -4.9912
2.7896 | -34.0764 | -5.0019
2.7958 | -34.0767 | -5.0127
2.8019 | -34.077 | -5.0233
2.808 | -34.0773 | -5.0344
2.8141 | -34.0776 | -5.0458
2.8202 | -34.0779 | -5.0568
2.8263 | -34.0782 | -5.068
2.8324 | -34.0785 | -5.0796
2.8385 | -34.0789 | -5.0906
2.8446 | -34.0793 | -5.1016
2.8507 | -34.0796 | -5.1125
2.8568 | -34.08 | -5.1232
2.8629 | -34.0802 | -5.1338
2.869 | -34.0806 | -5.1452
2.8751 | -34.0809 | -5.1561
2.8812 | -34.0813 | -5.1678
2.8873 | -34.0816 | -5.1788
2.8934 | -34.0819 | -5.1901
2.8995 | -34.0822 | -5.2014
2.9056 | -34.0825 | -5.2127
2.9117 | -34.0828 | -5.2241
2.9178 | -34.0832 | -5.2351
2.9239 | -34.0836 | -5.2469
2.93 | -34.084 | -5.2586
2.9361 | -34.0842 | -5.2698
2.9423 | -34.0846 | -5.2807
2.9484 | -34.085 | -5.2921
2.9545 | -34.0854 | -5.3034
2.9606 | -34.0858 | -5.3151
2.9667 | -34.0861 | -5.3264
2.9728 | -34.0864 | -5.3379
2.9789 | -34.0869 | -5.3491
2.985 | -34.0873 | -5.3606
2.9911 | -34.0876 | -5.3721
2.9972 | -34.0881 | -5.3836
3.0033 | -34.0885 | -5.3953
3.0094 | -34.0888 | -5.4066
3.0155 | -34.0892 | -5.4179
3.0216 | -34.0896 | -5.4292
3.0277 | -34.0899 | -5.4412
3.0338 | -34.0903 | -5.4527
3.0399 | -34.0907 | -5.4643
3.046 | -34.0909 | -5.4761
3.0521 | -34.0913 | -5.4873
3.0582 | -34.0915 | -5.4994
3.0643 | -34.0919 | -5.5111
3.0704 | -34.0922 | -5.5232
3.0765 | -34.0924 | -5.5346
3.0827 | -34.0927 | -5.5459
3.0888 | -34.093 | -5.5573
3.0949 | -34.0934 | -5.5687
3.101 | -34.0937 | -5.5804
3.1071 | -34.094 | -5.5924
3.1132 | -34.0944 | -5.6042
3.1193 | -34.0948 | -5.6158
3.1254 | -34.0951 | -5.6272
3.1315 | -34.0955 | -5.6393
3.1376 | -34.0957 | -5.651
3.1437 | -34.0961 | -5.6627
3.1498 | -34.0964 | -5.6749
3.1559 | -34.0967 | -5.6866
3.162 | -34.0969 | -5.6987
3.1681 | -34.0972 | -5.7107
3.1742 | -34.0975 | -5.7223
3.1803 | -34.0978 | -5.7341
3.1864 | -34.0981 | -5.7455
3.1925 | -34.0984 | -5.7571
3.1986 | -34.0987 | -5.7689
3.2047 | -34.0989 | -5.7811
3.2108 | -34.0991 | -5.7924
3.2169 | -34.0994 | -5.8042
3.223 | -34.0996 | -5.8159
3.2292 | -34.0999 | -5.8277
3.2353 | -34.1001 | -5.8386
3.2414 | -34.1003 | -5.85
3.2475 | -34.1005 | -5.8616
3.2536 | -34.1008 | -5.8731
3.2597 | -34.101 | -5.884
3.2658 | -34.1012 | -5.8951
3.2719 | -34.1014 | -5.906
3.278 | -34.1017 | -5.9176
3.2841 | -34.1019 | -5.9286
3.2902 | -34.1021 | -5.9401
3.2963 | -34.1024 | -5.9516
3.3024 | -34.1024 | -5.9627
3.3085 | -34.1027 | -5.9743
3.3146 | -34.1028 | -5.9856
3.3207 | -34.1031 | -5.9971
3.3268 | -34.1033 | -6.0086
3.3329 | -34.1034 | -6.0192
3.339 | -34.1036 | -6.0301
3.3451 | -34.1038 | -6.0413
3.3512 | -34.1039 | -6.052
3.3573 | -34.1041 | -6.0632
3.3634 | -34.1043 | -6.0745
3.3696 | -34.1045 | -6.0859
3.3757 | -34.1046 | -6.0968
3.3818 | -34.1048 | -6.1076
3.3879 | -34.105 | -6.1187
3.394 | -34.1051 | -6.1302
3.4001 | -34.1053 | -6.141
3.4062 | -34.1054 | -6.1519
3.4123 | -34.1055 | -6.1625
3.4184 | -34.1057 | -6.1738
3.4245 | -34.1059 | -6.1846
3.4306 | -34.106 | -6.1953
3.4367 | -34.1062 | -6.2059
3.4428 | -34.1063 | -6.2164
3.4489 | -34.1064 | -6.2273
3.455 | -34.1065 | -6.2383
3.4611 | -34.1067 | -6.2497
3.4672 | -34.1069 | -6.2608
3.4733 | -34.107 | -6.2719
3.4794 | -34.1071 | -6.2828
3.4855 | -34.1073 | -6.2932
3.4916 | -34.1075 | -6.3036
3.4977 | -34.1076 | -6.3148
3.5038 | -34.1077 | -6.3257
3.5099 | -34.1078 | -6.336
3.5161 | -34.108 | -6.3465
3.5222 | -34.1081 | -6.357
3.5283 | -34.1081 | -6.3678
3.5344 | -34.1083 | -6.3787
3.5405 | -34.1084 | -6.3893
3.5466 | -34.1086 | -6.4003
3.5527 | -34.1086 | -6.4112
3.5588 | -34.1087 | -6.422
3.5649 | -34.1088 | -6.4327
3.571 | -34.1089 | -6.4434
3.5771 | -34.109 | -6.4545
3.5832 | -34.1091 | -6.4645
3.5893 | -34.1093 | -6.4751
3.5954 | -34.1095 | -6.486
3.6015 | -34.1096 | -6.497
3.6076 | -34.1098 | -6.5076
3.6137 | -34.1099 | -6.5177
3.6198 | -34.1102 | -6.528
3.6259 | -34.1103 | -6.539
3.632 | -34.1104 | -6.5497
3.6381 | -34.1106 | -6.5605
3.6442 | -34.1107 | -6.5711
3.6503 | -34.1109 | -6.5814
3.6565 | -34.111 | -6.5922
3.6626 | -34.1111 | -6.6027
3.6687 | -34.1114 | -6.6134
3.6748 | -34.1116 | -6.6238
3.6809 | -34.1118 | -6.6345
3.687 | -34.112 | -6.6456
3.6931 | -34.1122 | -6.6558
3.6992 | -34.1124 | -6.6664
3.7053 | -34.1125 | -6.677
3.7114 | -34.1127 | -6.6877
3.7175 | -34.1129 | -6.6984
3.7236 | -34.1131 | -6.7085
3.7297 | -34.1133 | -6.7193
3.7358 | -34.1134 | -6.7303
3.7419 | -34.1136 | -6.7411
3.748 | -34.1138 | -6.7516
3.7541 | -34.114 | -6.7623
3.7602 | -34.1142 | -6.7728
3.7663 | -34.1144 | -6.7833
3.7724 | -34.1146 | -6.794
3.7785 | -34.1149 | -6.8044
3.7846 | -34.115 | -6.8149
3.7907 | -34.1153 | -6.8258
3.7969 | -34.1155 | -6.8365
3.803 | -34.1157 | -6.8467
3.8091 | -34.1159 | -6.8572
3.8152 | -34.1162 | -6.8677
3.8213 | -34.1164 | -6.8786
3.8274 | -34.1166 | -6.889
3.8335 | -34.1169 | -6.8997
3.8396 | -34.1172 | -6.9104
3.8457 | -34.1174 | -6.921
3.8518 | -34.1178 | -6.9315
3.8579 | -34.1181 | -6.9421
3.864 | -34.1183 | -6.9528
3.8701 | -34.1186 | -6.9637
3.8762 | -34.1188 | -6.9742
3.8823 | -34.119 | -6.9849
3.8884 | -34.1193 | -6.9955
3.8945 | -34.1196 | -7.0064
3.9006 | -34.1199 | -7.0173
3.9067 | -34.1201 | -7.0284
3.9128 | -34.1203 | -7.0392
3.9189 | -34.1206 | -7.0497
3.925 | -34.1208 | -7.0604
3.9311 | -34.121 | -7.0706
3.9372 | -34.1213 | -7.0815
3.9434 | -34.1216 | -7.0922
3.9495 | -34.1219 | -7.103
3.9556 | -34.1222 | -7.1134
3.9617 | -34.1225 | -7.1237
3.9678 | -34.1228 | -7.1341
3.9739 | -34.123 | -7.1447
3.98 | -34.1233 | -7.1555
3.9861 | -34.1235 | -7.1661
3.9922 | -34.1237 | -7.1768
3.9983 | -34.124 | -7.1869
4.0044 | -34.1242 | -7.1972
4.0105 | -34.1245 | -7.2078
4.0166 | -34.1248 | -7.2185
4.0227 | -34.1251 | -7.2291
4.0288 | -34.1253 | -7.2401
4.0349 | -34.1255 | -7.2505
4.041 | -34.1257 | -7.261
4.0471 | -34.1259 | -7.2718
4.0532 | -34.1262 | -7.2821
4.0593 | -34.1265 | -7.2929
4.0654 | -34.1267 | -7.3033
4.0715 | -34.127 | -7.3143
4.0776 | -34.1273 | -7.3248
4.0838 | -34.1275 | -7.3359
4.0899 | -34.1278 | -7.3462
4.096 | -34.128 | -7.3568
4.1021 | -34.1283 | -7.3671
4.1082 | -34.1285 | -7.3775
4.1143 | -34.1287 | -7.3878
4.1204 | -34.129 | -7.3981
4.1265 | -34.1293 | -7.4084
4.1326 | -34.1295 | -7.4187
4.1387 | -34.1298 | -7.4292
4.1448 | -34.13 | -7.4397
4.1509 | -34.1302 | -7.4503
4.157 | -34.1304 | -7.4612
4.1631 | -34.1307 | -7.472
4.1692 | -34.1309 | -7.4823
4.1753 | -34.1312 | -7.4928
4.1814 | -34.1314 | -7.503
4.1875 | -34.1316 | -7.5137
4.1936 | -34.1319 | -7.5244
4.1997 | -34.1321 | -7.5345
4.2058 | -34.1323 | -7.5453
4.2119 | -34.1325 | -7.5556
4.218 | -34.1327 | -7.5658
4.2241 | -34.1329 | -7.5761
4.2303 | -34.1332 | -7.5865
4.2364 | -34.1334 | -7.5969
4.2425 | -34.1336 | -7.6074
4.2486 | -34.1338 | -7.618
4.2547 | -34.1339 | -7.6282
4.2608 | -34.1341 | -7.6388
4.2669 | -34.1343 | -7.649
4.273 | -34.1345 | -7.6595
4.2791 | -34.1348 | -7.6699
4.2852 | -34.135 | -7.6805
4.2913 | -34.1351 | -7.6907
4.2974 | -34.1353 | -7.7014
4.3035 | -34.1356 | -7.7121
4.3096 | -34.1357 | -7.7228
4.3157 | -34.136 | -7.733
4.3218 | -34.1361 | -7.7433
4.3279 | -34.1364 | -7.7538
4.334 | -34.1365 | -7.7641
4.3401 | -34.1367 | -7.7742
4.3462 | -34.137 | -7.7847
4.3523 | -34.1372 | -7.7946
4.3584 | -34.1374 | -7.8047
4.3645 | -34.1376 | -7.8149
4.3707 | -34.1378 | -7.825
4.3768 | -34.138 | -7.8353
4.3829 | -34.1382 | -7.846
4.389 | -34.1384 | -7.8563
4.3951 | -34.1386 | -7.8664
4.4012 | -34.1387 | -7.8771
4.4073 | -34.1389 | -7.8872
4.4134 | -34.1391 | -7.8974
4.4195 | -34.1394 | -7.9077
4.4256 | -34.1396 | -7.9178
4.4317 | -34.1398 | -7.928
4.4378 | -34.14 | -7.9379
4.4439 | -34.1402 | -7.9479
4.45 | -34.1403 | -7.9581
4.4561 | -34.1405 | -7.9683
4.4622 | -34.1406 | -7.9786
4.4683 | -34.1408 | -7.9888
4.4744 | -34.141 | -7.9994
4.4805 | -34.1412 | -8.0092
4.4866 | -34.1414 | -8.0194
4.4927 | -34.1417 | -8.0298
4.4988 | -34.1419 | -8.04
4.5049 | -34.1421 | -8.0498
4.511 | -34.1423 | -8.0597
4.5172 | -34.1425 | -8.07
4.5233 | -34.1427 | -8.0806
4.5294 | -34.1429 | -8.0906
4.5355 | -34.1431 | -8.1008
4.5416 | -34.1433 | -8.1109
4.5477 | -34.1436 | -8.1212
4.5538 | -34.1437 | -8.131
4.5599 | -34.144 | -8.1411
4.566 | -34.1442 | -8.1512
4.5721 | -34.1444 | -8.1614
4.5782 | -34.1447 | -8.1718
4.5843 | -34.1449 | -8.1822
4.5904 | -34.1451 | -8.1924
4.5965 | -34.1454 | -8.2024
4.6026 | -34.1456 | -8.213
4.6087 | -34.1458 | -8.2231
4.6148 | -34.146 | -8.2337
4.6209 | -34.1462 | -8.2438
4.627 | -34.1465 | -8.2542
4.6331 | -34.1466 | -8.2642
4.6392 | -34.1468 | -8.2745
4.6453 | -34.1471 | -8.2847
4.6514 | -34.1473 | -8.2949
4.6576 | -34.1475 | -8.3055
4.6637 | -34.1477 | -8.3154
4.6698 | -34.1478 | -8.3255
4.6759 | -34.1481 | -8.3361
4.682 | -34.1483 | -8.3462
4.6881 | -34.1484 | -8.3564
4.6942 | -34.1487 | -8.3666
4.7003 | -34.1489 | -8.3769
4.7064 | -34.1491 | -8.387
4.7125 | -34.1494 | -8.3975
4.7186 | -34.1496 | -8.4075
4.7247 | -34.1498 | -8.4178
4.7308 | -34.1501 | -8.428
4.7369 | -34.1503 | -8.4383
4.743 | -34.1505 | -8.449
4.7491 | -34.1507 | -8.4595
4.7552 | -34.1509 | -8.4699
4.7613 | -34.1511 | -8.4803
4.7674 | -34.1514 | -8.4904
4.7735 | -34.1516 | -8.5011
4.7796 | -34.1519 | -8.5116
4.7857 | -34.1521 | -8.522
4.7918 | -34.1523 | -8.532
4.7979 | -34.1526 | -8.5422
4.8041 | -34.1528 | -8.5525
4.8102 | -34.1531 | -8.5624
4.8163 | -34.1533 | -8.5726
4.8224 | -34.1535 | -8.5829
4.8285 | -34.1538 | -8.5929
4.8346 | -34.1541 | -8.603
4.8407 | -34.1543 | -8.6134
4.8468 | -34.1545 | -8.624
4.8529 | -34.1548 | -8.6346
4.859 | -34.155 | -8.6452
4.8651 | -34.1552 | -8.6553
4.8712 | -34.1555 | -8.6654
4.8773 | -34.1557 | -8.6758
4.8834 | -34.1559 | -8.686
4.8895 | -34.1562 | -8.6962
4.8956 | -34.1564 | -8.7065
4.9017 | -34.1566 | -8.7167
4.9078 | -34.1568 | -8.7273
4.9139 | -34.1571 | -8.7372
4.92 | -34.1574 | -8.7473
4.9261 | -34.1577 | -8.7572
4.9322 | -34.1579 | -8.7674
4.9383 | -34.1582 | -8.7777
4.9445 | -34.1584 | -8.788
4.9506 | -34.1587 | -8.7982
4.9567 | -34.1589 | -8.8084
4.9628 | -34.1591 | -8.8186
4.9689 | -34.1594 | -8.8293
4.975 | -34.1596 | -8.84
4.9811 | -34.1599 | -8.85
4.9872 | -34.1602 | -8.8601
4.9933 | -34.1605 | -8.8707
4.9994 | -34.1608 | -8.8814
5.0055 | -34.1611 | -8.8918
5.0116 | -34.1613 | -8.9024
5.0177 | -34.1617 | -8.913
5.0238 | -34.1619 | -8.9237
5.0299 | -34.1623 | -8.9343
5.036 | -34.1626 | -8.9447
5.0421 | -34.1628 | -8.9549
5.0482 | -34.1631 | -8.9653
5.0543 | -34.1633 | -8.9754
5.0604 | -34.1636 | -8.9857
5.0665 | -34.1639 | -8.9958
5.0726 | -34.1642 | -9.0064
5.0787 | -34.1645 | -9.0169
5.0848 | -34.1648 | -9.0275
5.091 | -34.1651 | -9.0377
5.0971 | -34.1654 | -9.0476
5.1032 | -34.1657 | -9.0584
5.1093 | -34.166 | -9.0687
5.1154 | -34.1662 | -9.0788
5.1215 | -34.1664 | -9.089
5.1276 | -34.1667 | -9.0992
5.1337 | -34.1671 | -9.1093
5.1398 | -34.1673 | -9.1195
5.1459 | -34.1676 | -9.1301
5.152 | -34.1679 | -9.1406
5.1581 | -34.1682 | -9.1511
5.1642 | -34.1685 | -9.1616
5.1703 | -34.1688 | -9.1718
5.1764 | -34.1691 | -9.182
5.1825 | -34.1694 | -9.1922
5.1886 | -34.1697 | -9.2022
5.1947 | -34.1699 | -9.2125
5.2008 | -34.1703 | -9.223
5.2069 | -34.1705 | -9.2335
5.213 | -34.1708 | -9.2436
5.2191 | -34.1711 | -9.254
5.2252 | -34.1714 | -9.2644
5.2314 | -34.1717 | -9.275
5.2375 | -34.172 | -9.2851
5.2436 | -34.1722 | -9.2954
5.2497 | -34.1725 | -9.3055
5.2558 | -34.1729 | -9.3162
5.2619 | -34.1731 | -9.3269
5.268 | -34.1735 | -9.3374
5.2741 | -34.1738 | -9.3478
5.2802 | -34.1741 | -9.3585
5.2863 | -34.1744 | -9.3688
5.2924 | -34.1747 | -9.3793
5.2985 | -34.1751 | -9.3898
5.3046 | -34.1753 | -9.4002
5.3107 | -34.1756 | -9.4107
5.3168 | -34.1759 | -9.4211
5.3229 | -34.1762 | -9.4315
5.329 | -34.1765 | -9.4419
5.3351 | -34.1768 | -9.452
5.3412 | -34.1771 | -9.4623
5.3473 | -34.1774 | -9.4728
5.3534 | -34.1777 | -9.4832
5.3595 | -34.178 | -9.4936
5.3656 | -34.1784 | -9.504
5.3717 | -34.1786 | -9.5146
5.3779 | -34.1789 | -9.5254
5.384 | -34.1792 | -9.5358
5.3901 | -34.1796 | -9.5464
5.3962 | -34.1798 | -9.5566
5.4023 | -34.1802 | -9.5672
5.4084 | -34.1805 | -9.5775
5.4145 | -34.1808 | -9.5876
5.4206 | -34.1811 | -9.5981
5.4267 | -34.1814 | -9.6088
5.4328 | -34.1817 | -9.6192
5.4389 | -34.182 | -9.6293
5.445 | -34.1823 | -9.6401
5.4511 | -34.1826 | -9.6503
5.4572 | -34.1829 | -9.6608
5.4633 | -34.1832 | -9.6712
5.4694 | -34.1835 | -9.6816
5.4755 | -34.1838 | -9.6918
5.4816 | -34.1841 | -9.7022
5.4877 | -34.1844 | -9.7125
5.4938 | -34.1847 | -9.7226
5.4999 | -34.185 | -9.7329
5.506 | -34.1854 | -9.7432
5.5121 | -34.1857 | -9.7535
5.5183 | -34.186 | -9.7635
5.5244 | -34.1862 | -9.7741
5.5305 | -34.1866 | -9.7844
5.5366 | -34.1869 | -9.7945
5.5427 | -34.1871 | -9.8049
5.5488 | -34.1874 | -9.8153
5.5549 | -34.1877 | -9.8257
5.561 | -34.188 | -9.8363
5.5671 | -34.1883 | -9.8465
5.5732 | -34.1886 | -9.8571
5.5793 | -34.189 | -9.8677
5.5854 | -34.1893 | -9.878
5.5915 | -34.1896 | -9.8886
5.5976 | -34.1899 | -9.899
5.6037 | -34.1902 | -9.9094
5.6098 | -34.1904 | -9.9198
5.6159 | -34.1908 | -9.9299
5.622 | -34.1911 | -9.9403
5.6281 | -34.1914 | -9.9507
5.6342 | -34.1917 | -9.9614
5.6403 | -34.192 | -9.9718
5.6464 | -34.1923 | -9.9822
5.6525 | -34.1926 | -9.9928
5.6586 | -34.1929 | -10.0028
5.6648 | -34.1932 | -10.0132
5.6709 | -34.1935 | -10.0237
5.677 | -34.1938 | -10.0341
5.6831 | -34.1941 | -10.0445
5.6892 | -34.1944 | -10.055
5.6953 | -34.1947 | -10.0653
5.7014 | -34.195 | -10.0757
5.7075 | -34.1953 | -10.0861
5.7136 | -34.1956 | -10.0964
5.7197 | -34.1959 | -10.1068
5.7258 | -34.1961 | -10.1169
5.7319 | -34.1964 | -10.1273
5.738 | -34.1968 | -10.1376
5.7441 | -34.1971 | -10.1479
5.7502 | -34.1974 | -10.158
5.7563 | -34.1977 | -10.1684
5.7624 | -34.198 | -10.1789
5.7685 | -34.1983 | -10.1892
5.7746 | -34.1986 | -10.1998
5.7807 | -34.1989 | -10.2102
5.7868 | -34.1991 | -10.2208
5.7929 | -34.1994 | -10.231
5.799 | -34.1997 | -10.2412
5.8052 | -34.2 | -10.2514
5.8113 | -34.2003 | -10.2616
5.8174 | -34.2006 | -10.2718
5.8235 | -34.2009 | -10.2821
5.8296 | -34.2012 | -10.2926
5.8357 | -34.2014 | -10.3028
5.8418 | -34.2017 | -10.3132
5.8479 | -34.202 | -10.3237
5.854 | -34.2023 | -10.3342
5.8601 | -34.2026 | -10.3443
5.8662 | -34.2029 | -10.3548
5.8723 | -34.2032 | -10.365
5.8784 | -34.2035 | -10.3752
5.8845 | -34.2038 | -10.3854
5.8906 | -34.204 | -10.3958
5.8967 | -34.2043 | -10.4059
5.9028 | -34.2047 | -10.4161
5.9089 | -34.205 | -10.4262
5.915 | -34.2053 | -10.4363
5.9211 | -34.2056 | -10.4466
5.9272 | -34.2059 | -10.4566
5.9333 | -34.2062 | -10.4667
5.9394 | -34.2065 | -10.477
5.9455 | -34.2068 | -10.4871
5.9517 | -34.2072 | -10.4972
5.9578 | -34.2075 | -10.5076
5.9639 | -34.2078 | -10.5175
5.97 | -34.2082 | -10.5279
5.9761 | -34.2085 | -10.538
5.9822 | -34.2088 | -10.5485
5.9883 | -34.2091 | -10.5587
5.9944 | -34.2094 | -10.5691
6.0005 | -34.2097 | -10.5795
6.0066 | -34.21 | -10.5896
6.0127 | -34.2103 | -10.5995
6.0188 | -34.2106 | -10.61
6.0249 | -34.2109 | -10.6204
6.031 | -34.2112 | -10.6304
6.0371 | -34.2115 | -10.6406
6.0432 | -34.2118 | -10.6508
6.0493 | -34.2121 | -10.6609
6.0554 | -34.2125 | -10.671
6.0615 | -34.2128 | -10.6813
6.0676 | -34.2131 | -10.6916
6.0737 | -34.2134 | -10.702
6.0798 | -34.2137 | -10.7122
6.0859 | -34.2141 | -10.7224
6.0921 | -34.2143 | -10.7323
6.0982 | -34.2146 | -10.7426
6.1043 | -34.215 | -10.7527
6.1104 | -34.2153 | -10.7631
6.1165 | -34.2156 | -10.7729
6.1226 | -34.216 | -10.783
6.1287 | -34.2163 | -10.7932
6.1348 | -34.2166 | -10.8034
6.1409 | -34.2169 | -10.8135
6.147 | -34.2172 | -10.8236
6.1531 | -34.2175 | -10.8338
6.1592 | -34.2178 | -10.8439
6.1653 | -34.2181 | -10.8539
6.1714 | -34.2184 | -10.8642
6.1775 | -34.2187 | -10.8745
6.1836 | -34.219 | -10.8847
6.1897 | -34.2194 | -10.8949
6.1958 | -34.2197 | -10.9054
6.2019 | -34.22 | -10.9156
6.208 | -34.2203 | -10.9258
6.2141 | -34.2207 | -10.9359
6.2202 | -34.221 | -10.9462
6.2263 | -34.2213 | -10.9564
6.2325 | -34.2216 | -10.9663
6.2386 | -34.2219 | -10.9762
6.2447 | -34.2223 | -10.9866
6.2508 | -34.2226 | -10.9968
6.2569 | -34.2229 | -11.0069
6.263 | -34.2232 | -11.017
6.2691 | -34.2235 | -11.0271
6.2752 | -34.2238 | -11.0374
6.2813 | -34.2241 | -11.0475
6.2874 | -34.2244 | -11.0576
6.2935 | -34.2247 | -11.0678
6.2996 | -34.225 | -11.0779
6.3057 | -34.2253 | -11.0881
6.3118 | -34.2256 | -11.0984
6.3179 | -34.2259 | -11.1086
6.324 | -34.2262 | -11.1188
6.3301 | -34.2266 | -11.129
6.3362 | -34.2269 | -11.1391
6.3423 | -34.2272 | -11.1495
6.3484 | -34.2276 | -11.1597
6.3545 | -34.2278 | -11.1696
6.3606 | -34.2281 | -11.1795
6.3667 | -34.2284 | -11.1896
6.3728 | -34.2288 | -11.1997
6.379 | -34.2291 | -11.2099
6.3851 | -34.2294 | -11.2199
6.3912 | -34.2297 | -11.23
6.3973 | -34.23 | -11.2402
6.4034 | -34.2303 | -11.2502
6.4095 | -34.2306 | -11.2603
6.4156 | -34.2309 | -11.2705
6.4217 | -34.2312 | -11.2806
6.4278 | -34.2315 | -11.2908
6.4339 | -34.2318 | -11.3009
6.44 | -34.2322 | -11.3109
6.4461 | -34.2325 | -11.3212
6.4522 | -34.2328 | -11.3312
6.4583 | -34.233 | -11.3414
6.4644 | -34.2333 | -11.3517
6.4705 | -34.2336 | -11.3616
6.4766 | -34.234 | -11.3717
6.4827 | -34.2343 | -11.3819
6.4888 | -34.2345 | -11.3919
6.4949 | -34.2348 | -11.4018
6.501 | -34.2352 | -11.4121
6.5071 | -34.2355 | -11.422
6.5132 | -34.2358 | -11.4322
6.5194 | -34.2361 | -11.4423
6.5255 | -34.2364 | -11.4522
6.5316 | -34.2367 | -11.4625
6.5377 | -34.237 | -11.4726
6.5438 | -34.2373 | -11.4826
6.5499 | -34.2376 | -11.4926
6.556 | -34.2379 | -11.5027
6.5621 | -34.2382 | -11.513
6.5682 | -34.2385 | -11.5231
6.5743 | -34.2389 | -11.5332
6.5804 | -34.2392 | -11.5435
6.5865 | -34.2395 | -11.5536
6.5926 | -34.2398 | -11.5636
6.5987 | -34.2401 | -11.5735
6.6048 | -34.2404 | -11.5837
6.6109 | -34.2407 | -11.5939
6.617 | -34.241 | -11.6041
6.6231 | -34.2413 | -11.6142
6.6292 | -34.2416 | -11.6244
6.6353 | -34.2419 | -11.6348
6.6414 | -34.2423 | -11.6448
6.6475 | -34.2426 | -11.655
6.6536 | -34.2429 | -11.6651
6.6597 | -34.2432 | -11.6749
6.6659 | -34.2435 | -11.685
6.672 | -34.2438 | -11.6951
6.6781 | -34.2441 | -11.7052
6.6842 | -34.2444 | -11.7152
6.6903 | -34.2448 | -11.7252
6.6964 | -34.2451 | -11.7352
6.7025 | -34.2454 | -11.7451
6.7086 | -34.2457 | -11.7552
6.7147 | -34.246 | -11.7654
6.7208 | -34.2463 | -11.7756
6.7269 | -34.2466 | -11.7856
6.733 | -34.2469 | -11.7955
6.7391 | -34.2472 | -11.8057
6.7452 | -34.2475 | -11.8155
6.7513 | -34.2478 | -11.8254
6.7574 | -34.2481 | -11.8355
6.7635 | -34.2485 | -11.8454
6.7696 | -34.2488 | -11.8554
6.7757 | -34.2491 | -11.8654
6.7818 | -34.2494 | -11.8755
6.7879 | -34.2498 | -11.8855
6.794 | -34.2501 | -11.8954
6.8001 | -34.2503 | -11.9054
6.8063 | -34.2507 | -11.9155
6.8124 | -34.251 | -11.9257
6.8185 | -34.2513 | -11.9359
6.8246 | -34.2515 | -11.9459
6.8307 | -34.2519 | -11.9559
6.8368 | -34.2522 | -11.9661
6.8429 | -34.2525 | -11.9763
6.849 | -34.2528 | -11.9863
6.8551 | -34.2532 | -11.9963
6.8612 | -34.2535 | -12.0063
6.8673 | -34.2538 | -12.0163
6.8734 | -34.2541 | -12.0265
6.8795 | -34.2544 | -12.0366
6.8856 | -34.2547 | -12.0466
6.8917 | -34.255 | -12.0567
6.8978 | -34.2553 | -12.0665
6.9039 | -34.2556 | -12.0763
6.91 | -34.256 | -12.0864
6.9161 | -34.2563 | -12.0962
6.9222 | -34.2566 | -12.1062
6.9283 | -34.257 | -12.116
6.9344 | -34.2573 | -12.1261
6.9405 | -34.2576 | -12.136
6.9466 | -34.258 | -12.1463
6.9528 | -34.2583 | -12.1565
6.9589 | -34.2586 | -12.1666
6.965 | -34.259 | -12.1767
6.9711 | -34.2593 | -12.1869
6.9772 | -34.2597 | -12.1968
6.9833 | -34.26 | -12.2068
6.9894 | -34.2603 | -12.2171
6.9955 | -34.2606 | -12.2269
7.0016 | -34.261 | -12.2369
7.0077 | -34.2613 | -12.2468
7.0138 | -34.2616 | -12.2568
7.0199 | -34.2619 | -12.2668
7.026 | -34.2622 | -12.2767
7.0321 | -34.2626 | -12.2865
7.0382 | -34.2629 | -12.2962
7.0443 | -34.2632 | -12.3063
7.0504 | -34.2636 | -12.3164
7.0565 | -34.2639 | -12.3264
7.0626 | -34.2642 | -12.3361
7.0687 | -34.2645 | -12.3457
7.0748 | -34.2649 | -12.3555
7.0809 | -34.2652 | -12.3654
7.087 | -34.2655 | -12.3755
7.0932 | -34.2658 | -12.3855
7.0993 | -34.2662 | -12.3954
7.1054 | -34.2665 | -12.4053
7.1115 | -34.2669 | -12.415
7.1176 | -34.2672 | -12.4249
7.1237 | -34.2675 | -12.4349
7.1298 | -34.2678 | -12.4447
7.1359 | -34.2681 | -12.4547
7.142 | -34.2685 | -12.4644
7.1481 | -34.2688 | -12.4742
7.1542 | -34.2691 | -12.4841
7.1603 | -34.2694 | -12.4941
7.1664 | -34.2698 | -12.5039
7.1725 | -34.2701 | -12.5138
7.1786 | -34.2705 | -12.5237
7.1847 | -34.2708 | -12.5337
7.1908 | -34.2711 | -12.5438
7.1969 | -34.2714 | -12.5539
7.203 | -34.2718 | -12.5636
7.2091 | -34.2721 | -12.5734
7.2152 | -34.2724 | -12.5833
7.2213 | -34.2728 | -12.593
7.2274 | -34.2731 | -12.603
7.2335 | -34.2734 | -12.6128
7.2397 | -34.2738 | -12.6226
7.2458 | -34.2742 | -12.6322
7.2519 | -34.2745 | -12.6421
7.258 | -34.2748 | -12.6522
7.2641 | -34.2751 | -12.662
7.2702 | -34.2755 | -12.6721
7.2763 | -34.2758 | -12.682
7.2824 | -34.2762 | -12.692
7.2885 | -34.2765 | -12.7018
7.2946 | -34.2768 | -12.7117
7.3007 | -34.2771 | -12.7217
7.3068 | -34.2774 | -12.7318
7.3129 | -34.2778 | -12.7417
7.319 | -34.2782 | -12.7516
7.3251 | -34.2785 | -12.7613
7.3312 | -34.2788 | -12.7712
7.3373 | -34.2792 | -12.7808
7.3434 | -34.2795 | -12.7906
7.3495 | -34.2799 | -12.8003
7.3556 | -34.2802 | -12.8102
7.3617 | -34.2806 | -12.8203
7.3678 | -34.2809 | -12.8301
7.3739 | -34.2812 | -12.8401
7.3801 | -34.2815 | -12.8502
7.3862 | -34.2819 | -12.86
7.3923 | -34.2822 | -12.87
7.3984 | -34.2826 | -12.8797
7.4045 | -34.2829 | -12.8895
7.4106 | -34.2833 | -12.8993
7.4167 | -34.2836 | -12.9093
7.4228 | -34.284 | -12.9192
7.4289 | -34.2843 | -12.929
7.435 | -34.2847 | -12.9386
7.4411 | -34.285 | -12.9485
7.4472 | -34.2853 | -12.9584
7.4533 | -34.2857 | -12.968
7.4594 | -34.2861 | -12.9781
7.4655 | -34.2864 | -12.9878
7.4716 | -34.2867 | -12.9978
7.4777 | -34.2871 | -13.0076
7.4838 | -34.2874 | -13.0173
7.4899 | -34.2877 | -13.027
7.496 | -34.2881 | -13.0366
7.5021 | -34.2884 | -13.0464
7.5082 | -34.2888 | -13.0563
7.5143 | -34.2891 | -13.0662
7.5204 | -34.2894 | -13.0761
7.5266 | -34.2898 | -13.0858
7.5327 | -34.2901 | -13.0955
7.5388 | -34.2904 | -13.1056
7.5449 | -34.2908 | -13.1156
7.551 | -34.2912 | -13.1253
7.5571 | -34.2915 | -13.1353
7.5632 | -34.2919 | -13.1451
7.5693 | -34.2922 | -13.1548
7.5754 | -34.2926 | -13.1645
7.5815 | -34.2929 | -13.1743
7.5876 | -34.2932 | -13.1844
7.5937 | -34.2936 | -13.1943
7.5998 | -34.294 | -13.204
7.6059 | -34.2943 | -13.2141
7.612 | -34.2947 | -13.2239
7.6181 | -34.295 | -13.2336
7.6242 | -34.2954 | -13.2437
7.6303 | -34.2957 | -13.2537
7.6364 | -34.296 | -13.2636
7.6425 | -34.2964 | -13.2735
7.6486 | -34.2967 | -13.2833
7.6547 | -34.2971 | -13.2931
7.6608 | -34.2974 | -13.3028
7.667 | -34.2977 | -13.3127
7.6731 | -34.2981 | -13.3226
7.6792 | -34.2985 | -13.3327
7.6853 | -34.2988 | -13.3426
7.6914 | -34.2991 | -13.3526
7.6975 | -34.2995 | -13.3625
7.7036 | -34.2999 | -13.3724
7.7097 | -34.3002 | -13.3824
7.7158 | -34.3006 | -13.3923
7.7219 | -34.3009 | -13.4019
7.728 | -34.3013 | -13.4118
7.7341 | -34.3016 | -13.4217
7.7402 | -34.302 | -13.4316
7.7463 | -34.3023 | -13.4415
7.7524 | -34.3027 | -13.4516
7.7585 | -34.303 | -13.4615
7.7646 | -34.3033 | -13.4713
7.7707 | -34.3037 | -13.4812
7.7768 | -34.304 | -13.4912
7.7829 | -34.3044 | -13.5014
7.789 | -34.3047 | -13.5115
7.7951 | -34.305 | -13.5213
7.8012 | -34.3054 | -13.5312
7.8073 | -34.3058 | -13.5412
7.8135 | -34.3061 | -13.5511
7.8196 | -34.3065 | -13.5611
7.8257 | -34.3068 | -13.5712
7.8318 | -34.3072 | -13.5811
7.8379 | -34.3075 | -13.5909
7.844 | -34.3079 | -13.6007
7.8501 | -34.3082 | -13.6108
7.8562 | -34.3085 | -13.621
7.8623 | -34.3089 | -13.6308
7.8684 | -34.3092 | -13.6407
7.8745 | -34.3096 | -13.6507
7.8806 | -34.3099 | -13.6607
7.8867 | -34.3103 | -13.6704
7.8928 | -34.3106 | -13.6801
7.8989 | -34.311 | -13.69
7.905 | -34.3113 | -13.7
7.9111 | -34.3117 | -13.7099
7.9172 | -34.312 | -13.7198
7.9233 | -34.3123 | -13.7298
7.9294 | -34.3127 | -13.7398
7.9355 | -34.313 | -13.7495
7.9416 | -34.3134 | -13.7595
7.9477 | -34.3137 | -13.7692
7.9539 | -34.314 | -13.779
7.96 | -34.3144 | -13.789
7.9661 | -34.3147 | -13.799
7.9722 | -34.3151 | -13.8087
7.9783 | -34.3154 | -13.8188
7.9844 | -34.3158 | -13.8286
7.9905 | -34.3161 | -13.8387
7.9966 | -34.3165 | -13.8487
8.0027 | -34.3168 | -13.8586
8.0088 | -34.3172 | -13.8685
8.0149 | -34.3175 | -13.8784
8.021 | -34.3178 | -13.8884
8.0271 | -34.3182 | -13.8985
8.0332 | -34.3185 | -13.9083
8.0393 | -34.3189 | -13.9183
8.0454 | -34.3192 | -13.9284
8.0515 | -34.3196 | -13.9384
8.0576 | -34.32 | -13.9485
8.0637 | -34.3203 | -13.9583
8.0698 | -34.3206 | -13.9681
8.0759 | -34.321 | -13.9779
8.082 | -34.3213 | -13.9877
8.0881 | -34.3217 | -13.9976
8.0942 | -34.322 | -14.0076
8.1004 | -34.3224 | -14.0176
8.1065 | -34.3227 | -14.0274
8.1126 | -34.3231 | -14.0374
8.1187 | -34.3234 | -14.0474
8.1248 | -34.3237 | -14.0571
8.1309 | -34.3241 | -14.0669
8.137 | -34.3245 | -14.0765
8.1431 | -34.3248 | -14.0867
8.1492 | -34.3252 | -14.0965
8.1553 | -34.3255 | -14.1065
8.1614 | -34.3258 | -14.1166
8.1675 | -34.3262 | -14.1265
8.1736 | -34.3265 | -14.1366
8.1797 | -34.3268 | -14.1465
8.1858 | -34.3272 | -14.1565
8.1919 | -34.3275 | -14.1664
8.198 | -34.3279 | -14.1761
8.2041 | -34.3283 | -14.1861
8.2102 | -34.3286 | -14.196
8.2163 | -34.3289 | -14.2058
8.2224 | -34.3293 | -14.2157
8.2285 | -34.3296 | -14.2254
8.2346 | -34.33 | -14.2352
8.2408 | -34.3303 | -14.2449
8.2469 | -34.3307 | -14.2549
8.253 | -34.331 | -14.2647
8.2591 | -34.3314 | -14.2746
8.2652 | -34.3318 | -14.2845
8.2713 | -34.3321 | -14.2945
8.2774 | -34.3325 | -14.3044
8.2835 | -34.3328 | -14.3146
8.2896 | -34.3332 | -14.3243
8.2957 | -34.3335 | -14.3343
8.3018 | -34.3339 | -14.3442
8.3079 | -34.3342 | -14.354
8.314 | -34.3346 | -14.3639
8.3201 | -34.335 | -14.3737
8.3262 | -34.3353 | -14.3836
8.3323 | -34.3356 | -14.3934
8.3384 | -34.336 | -14.4035
8.3445 | -34.3363 | -14.4132
8.3506 | -34.3367 | -14.4232
8.3567 | -34.3371 | -14.4332
8.3628 | -34.3374 | -14.4431
8.3689 | -34.3377 | -14.4529
8.375 | -34.3381 | -14.4628
8.3812 | -34.3384 | -14.4725
8.3873 | -34.3388 | -14.4824
8.3934 | -34.3391 | -14.4924
8.3995 | -34.3395 | -14.5021
8.4056 | -34.3398 | -14.512
8.4117 | -34.3402 | -14.5219
8.4178 | -34.3405 | -14.5318
8.4239 | -34.3408 | -14.5415
8.43 | -34.3412 | -14.5515
8.4361 | -34.3415 | -14.5614
8.4422 | -34.3419 | -14.5712
8.4483 | -34.3422 | -14.581
8.4544 | -34.3426 | -14.5908
8.4605 | -34.3429 | -14.6006
8.4666 | -34.3433 | -14.6106
8.4727 | -34.3436 | -14.6205
8.4788 | -34.344 | -14.6304
8.4849 | -34.3443 | -14.6402
8.491 | -34.3446 | -14.6501
8.4971 | -34.345 | -14.6598
8.5032 | -34.3453 | -14.6695
8.5093 | -34.3457 | -14.6792
8.5154 | -34.3461 | -14.6891
8.5215 | -34.3465 | -14.6988
8.5277 | -34.3468 | -14.7087
8.5338 | -34.3471 | -14.7185
8.5399 | -34.3475 | -14.7283
8.546 | -34.3479 | -14.738
8.5521 | -34.3482 | -14.7478
8.5582 | -34.3485 | -14.7577
8.5643 | -34.3489 | -14.7673
8.5704 | -34.3492 | -14.7769
8.5765 | -34.3496 | -14.7868
8.5826 | -34.35 | -14.7963
8.5887 | -34.3503 | -14.806
8.5948 | -34.3507 | -14.8159
8.6009 | -34.3511 | -14.8255
8.607 | -34.3514 | -14.8354
8.6131 | -34.3518 | -14.8452
8.6192 | -34.3522 | -14.8549
8.6253 | -34.3525 | -14.8646
8.6314 | -34.3529 | -14.8744
8.6375 | -34.3532 | -14.8842
8.6436 | -34.3536 | -14.894
8.6497 | -34.3539 | -14.9037
8.6558 | -34.3543 | -14.9135
8.6619 | -34.3546 | -14.9233
8.6681 | -34.355 | -14.9333
8.6742 | -34.3553 | -14.943
8.6803 | -34.3557 | -14.9527
8.6864 | -34.3561 | -14.9622
8.6925 | -34.3564 | -14.9718
8.6986 | -34.3568 | -14.9816
8.7047 | -34.3571 | -14.9916
8.7108 | -34.3575 | -15.0013
8.7169 | -34.3578 | -15.0111
8.723 | -34.3582 | -15.0208
8.7291 | -34.3585 | -15.0304
8.7352 | -34.3589 | -15.0403
8.7413 | -34.3592 | -15.0501
8.7474 | -34.3596 | -15.0597
8.7535 | -34.36 | -15.0695
8.7596 | -34.3604 | -15.0793
8.7657 | -34.3607 | -15.0891
8.7718 | -34.361 | -15.0989
8.7779 | -34.3614 | -15.1088
8.784 | -34.3618 | -15.1186
8.7901 | -34.3622 | -15.1284
8.7962 | -34.3625 | -15.1379
8.8023 | -34.3629 | -15.1476
8.8084 | -34.3632 | -15.1573
8.8146 | -34.3636 | -15.167
8.8207 | -34.3639 | -15.1767
8.8268 | -34.3643 | -15.1862
8.8329 | -34.3647 | -15.1959
8.839 | -34.3651 | -15.2057
8.8451 | -34.3654 | -15.2153
8.8512 | -34.3658 | -15.2249
8.8573 | -34.3661 | -15.2347
8.8634 | -34.3665 | -15.2442
8.8695 | -34.3669 | -15.2541
8.8756 | -34.3672 | -15.2639
8.8817 | -34.3676 | -15.2735
8.8878 | -34.368 | -15.2834
8.8939 | -34.3683 | -15.2931
8.9 | -34.3687 | -15.3028
8.9061 | -34.369 | -15.3124
8.9122 | -34.3694 | -15.3222
8.9183 | -34.3698 | -15.3319
8.9244 | -34.3702 | -15.3418
8.9305 | -34.3706 | -15.3512
8.9366 | -34.371 | -15.3609
8.9427 | -34.3713 | -15.3707
8.9488 | -34.3717 | -15.3804
8.955 | -34.3721 | -15.39
8.9611 | -34.3724 | -15.3999
8.9672 | -34.3729 | -15.4095
8.9733 | -34.3732 | -15.419
8.9794 | -34.3736 | -15.4286
8.9855 | -34.374 | -15.4384
8.9916 | -34.3744 | -15.4482
8.9977 | -34.3747 | -15.4578
9.0038 | -34.3752 | -15.4674
9.0099 | -34.3755 | -15.4771
9.016 | -34.3759 | -15.4868
9.0221 | -34.3762 | -15.4963
9.0282 | -34.3766 | -15.5058
9.0343 | -34.377 | -15.5154
9.0404 | -34.3774 | -15.5251
9.0465 | -34.3778 | -15.535
9.0526 | -34.3782 | -15.5445
9.0587 | -34.3786 | -15.5542
9.0648 | -34.379 | -15.5638
9.0709 | -34.3793 | -15.5734
9.077 | -34.3797 | -15.5832
9.0831 | -34.3801 | -15.5929
9.0892 | -34.3805 | -15.6026
9.0953 | -34.3809 | -15.6124
9.1015 | -34.3813 | -15.622
9.1076 | -34.3817 | -15.6316
9.1137 | -34.382 | -15.6412
9.1198 | -34.3825 | -15.6508
9.1259 | -34.3828 | -15.6604
9.132 | -34.3832 | -15.6702
9.1381 | -34.3836 | -15.6798
9.1442 | -34.384 | -15.6891
9.1503 | -34.3844 | -15.6989
9.1564 | -34.3848 | -15.7085
9.1625 | -34.3852 | -15.7182
9.1686 | -34.3856 | -15.7278
9.1747 | -34.386 | -15.7375
9.1808 | -34.3864 | -15.7471
9.1869 | -34.3868 | -15.7568
9.193 | -34.3871 | -15.7665
9.1991 | -34.3875 | -15.7762
9.2052 | -34.3879 | -15.7858
9.2113 | -34.3883 | -15.7955
9.2174 | -34.3888 | -15.8051
9.2235 | -34.3892 | -15.8148
9.2296 | -34.3896 | -15.8244
9.2357 | -34.39 | -15.8342
9.2419 | -34.3904 | -15.8438
9.248 | -34.3908 | -15.8534
9.2541 | -34.3912 | -15.8631
9.2602 | -34.3916 | -15.8729
9.2663 | -34.392 | -15.8823
9.2724 | -34.3924 | -15.8918
9.2785 | -34.3927 | -15.9014
9.2846 | -34.3932 | -15.9109
9.2907 | -34.3936 | -15.9205
9.2968 | -34.394 | -15.9303
9.3029 | -34.3944 | -15.94
9.309 | -34.3948 | -15.9496
9.3151 | -34.3951 | -15.9591
9.3212 | -34.3956 | -15.9686
9.3273 | -34.396 | -15.9782
9.3334 | -34.3964 | -15.9876
9.3395 | -34.3968 | -15.9972
9.3456 | -34.3971 | -16.0066
9.3517 | -34.3975 | -16.0163
9.3578 | -34.3979 | -16.0259
9.3639 | -34.3983 | -16.0354
9.37 | -34.3987 | -16.045
9.3761 | -34.3991 | -16.0548
9.3822 | -34.3996 | -16.0645
9.3884 | -34.4 | -16.0738
9.3945 | -34.4004 | -16.0835
9.4006 | -34.4008 | -16.0931
9.4067 | -34.4012 | -16.1027
9.4128 | -34.4016 | -16.1123
9.4189 | -34.402 | -16.1221
9.425 | -34.4024 | -16.1316
9.4311 | -34.4028 | -16.1412
9.4372 | -34.4033 | -16.1507
9.4433 | -34.4037 | -16.1603
9.4494 | -34.4041 | -16.1698
9.4555 | -34.4045 | -16.1795
9.4616 | -34.4049 | -16.1891
9.4677 | -34.4053 | -16.1987
9.4738 | -34.4057 | -16.2083
9.4799 | -34.4062 | -16.2179
9.486 | -34.4066 | -16.2275
9.4921 | -34.407 | -16.237
9.4982 | -34.4074 | -16.2464
9.5043 | -34.4078 | -16.256
9.5104 | -34.4082 | -16.2654
9.5165 | -34.4086 | -16.275
9.5226 | -34.409 | -16.2848
9.5288 | -34.4094 | -16.2942
9.5349 | -34.4098 | -16.3038
9.541 | -34.4102 | -16.3132
9.5471 | -34.4106 | -16.3227
9.5532 | -34.411 | -16.3324
9.5593 | -34.4115 | -16.3421
9.5654 | -34.4119 | -16.3516
9.5715 | -34.4123 | -16.3611
9.5776 | -34.4127 | -16.3708
9.5837 | -34.413 | -16.3803
9.5898 | -34.4134 | -16.39
9.5959 | -34.4139 | -16.3996
9.602 | -34.4143 | -16.4089
9.6081 | -34.4147 | -16.4185
9.6142 | -34.4151 | -16.428
9.6203 | -34.4155 | -16.4376
9.6264 | -34.4159 | -16.4471
9.6325 | -34.4163 | -16.4568
9.6386 | -34.4167 | -16.4662
9.6447 | -34.4171 | -16.4758
9.6508 | -34.4175 | -16.4853
9.6569 | -34.4179 | -16.4949
9.663 | -34.4183 | -16.5045
9.6691 | -34.4188 | -16.5141
9.6753 | -34.4192 | -16.5236
9.6814 | -34.4196 | -16.5329
9.6875 | -34.42 | -16.5423
9.6936 | -34.4204 | -16.5517
9.6997 | -34.4208 | -16.5613
9.7058 | -34.4213 | -16.5708
9.7119 | -34.4217 | -16.5805
9.718 | -34.4221 | -16.5901
9.7241 | -34.4225 | -16.5996
9.7302 | -34.4229 | -16.6091
9.7363 | -34.4233 | -16.6185
9.7424 | -34.4238 | -16.6279
9.7485 | -34.4242 | -16.6375
9.7546 | -34.4246 | -16.6469
9.7607 | -34.425 | -16.6566
9.7668 | -34.4254 | -16.6662
9.7729 | -34.4258 | -16.6756
9.779 | -34.4262 | -16.6849
9.7851 | -34.4267 | -16.6946
9.7912 | -34.4271 | -16.704
9.7973 | -34.4275 | -16.7136
9.8034 | -34.4279 | -16.7229
9.8095 | -34.4283 | -16.7326
9.8157 | -34.4288 | -16.7421
9.8218 | -34.4292 | -16.7518
9.8279 | -34.4296 | -16.7612
9.834 | -34.43 | -16.7708
9.8401 | -34.4304 | -16.7803
9.8462 | -34.4309 | -16.7896
9.8523 | -34.4312 | -16.7991
9.8584 | -34.4316 | -16.8086
9.8645 | -34.432 | -16.8183
9.8706 | -34.4324 | -16.8279
9.8767 | -34.4328 | -16.8373
9.8828 | -34.4332 | -16.8467
9.8889 | -34.4337 | -16.856
9.895 | -34.4341 | -16.8656
9.9011 | -34.4345 | -16.8751
9.9072 | -34.4349 | -16.8846
9.9133 | -34.4353 | -16.894
9.9194 | -34.4357 | -16.9034
9.9255 | -34.4361 | -16.9128
9.9316 | -34.4365 | -16.9223
9.9377 | -34.4369 | -16.9318
9.9438 | -34.4373 | -16.9413
9.9499 | -34.4377 | -16.951
9.956 | -34.4382 | -16.9605
9.9622 | -34.4386 | -16.9701
9.9683 | -34.439 | -16.9796
9.9744 | -34.4394 | -16.9893
9.9805 | -34.4398 | -16.9987
9.9866 | -34.4402 | -17.0081
9.9927 | -34.4406 | -17.0174
9.9988 | -34.441 | -17.027
10.0049 | -34.4414 | -17.0364
10.011 | -34.4418 | -17.046
10.0171 | -34.4422 | -17.0554
10.0232 | -34.4427 | -17.0649
10.0293 | -34.4431 | -17.0745
10.0354 | -34.4435 | -17.0839
10.0415 | -34.4439 | -17.0934
10.0476 | -34.4443 | -17.103
10.0537 | -34.4447 | -17.1125
10.0598 | -34.4451 | -17.1219
10.0659 | -34.4455 | -17.1315
10.072 | -34.4459 | -17.1409
10.0781 | -34.4464 | -17.1503
10.0842 | -34.4468 | -17.1598
10.0903 | -34.4472 | -17.1692
10.0964 | -34.4476 | -17.1784
10.1026 | -34.448 | -17.1879
10.1087 | -34.4484 | -17.1973
10.1148 | -34.4488 | -17.2069
10.1209 | -34.4492 | -17.2162
10.127 | -34.4496 | -17.2256
10.1331 | -34.45 | -17.2349
10.1392 | -34.4504 | -17.2442
10.1453 | -34.4508 | -17.2536
10.1514 | -34.4512 | -17.263
10.1575 | -34.4516 | -17.2726
10.1636 | -34.452 | -17.2817
10.1697 | -34.4524 | -17.2913
10.1758 | -34.4529 | -17.3006
10.1819 | -34.4533 | -17.3099
10.188 | -34.4537 | -17.3194
10.1941 | -34.4541 | -17.3289
10.2002 | -34.4545 | -17.3383
10.2063 | -34.4549 | -17.3476
10.2124 | -34.4553 | -17.3571
10.2185 | -34.4557 | -17.3666
10.2246 | -34.4561 | -17.3759
10.2307 | -34.4565 | -17.3854
10.2368 | -34.4569 | -17.3945
10.2429 | -34.4573 | -17.4039
10.2491 | -34.4577 | -17.4134
10.2552 | -34.4581 | -17.4229
10.2613 | -34.4585 | -17.4324
10.2674 | -34.4589 | -17.4418
10.2735 | -34.4593 | -17.4514
10.2796 | -34.4598 | -17.4608
10.2857 | -34.4602 | -17.4704
10.2918 | -34.4606 | -17.4799
10.2979 | -34.461 | -17.4893
10.304 | -34.4614 | -17.4986
10.3101 | -34.4618 | -17.508
10.3162 | -34.4622 | -17.5175
10.3223 | -34.4626 | -17.5267
10.3284 | -34.463 | -17.5361
10.3345 | -34.4634 | -17.5453
10.3406 | -34.4638 | -17.5547
10.3467 | -34.4642 | -17.564
10.3528 | -34.4646 | -17.5733
10.3589 | -34.465 | -17.5827
10.365 | -34.4654 | -17.5922
10.3711 | -34.4658 | -17.6016
10.3772 | -34.4663 | -17.6111
10.3833 | -34.4667 | -17.6204
10.3895 | -34.4671 | -17.6298
10.3956 | -34.4675 | -17.6392
10.4017 | -34.4679 | -17.6485
10.4078 | -34.4683 | -17.6578
10.4139 | -34.4687 | -17.6672
10.42 | -34.4691 | -17.6766
10.4261 | -34.4695 | -17.686
10.4322 | -34.4699 | -17.6954
10.4383 | -34.4703 | -17.7047
10.4444 | -34.4707 | -17.7139
10.4505 | -34.4711 | -17.7233
10.4566 | -34.4715 | -17.7326
10.4627 | -34.4719 | -17.7418
10.4688 | -34.4723 | -17.751
10.4749 | -34.4727 | -17.7605
10.481 | -34.4731 | -17.7699
10.4871 | -34.4735 | -17.7792
10.4932 | -34.474 | -17.7884
10.4993 | -34.4744 | -17.7977
10.5054 | -34.4747 | -17.807
10.5115 | -34.4752 | -17.8164
10.5176 | -34.4756 | -17.8256
10.5237 | -34.476 | -17.835
10.5298 | -34.4764 | -17.8443
10.536 | -34.4768 | -17.8538
10.5421 | -34.4772 | -17.863
10.5482 | -34.4776 | -17.8724
10.5543 | -34.4781 | -17.8816
10.5604 | -34.4784 | -17.8908
10.5665 | -34.4789 | -17.9001
10.5726 | -34.4793 | -17.9094
10.5787 | -34.4797 | -17.9186
10.5848 | -34.4801 | -17.9276
10.5909 | -34.4805 | -17.937
10.597 | -34.4809 | -17.9464
10.6031 | -34.4813 | -17.9559
10.6092 | -34.4817 | -17.9651
10.6153 | -34.4821 | -17.9746
10.6214 | -34.4825 | -17.9837
10.6275 | -34.4829 | -17.9931
10.6336 | -34.4833 | -18.0024
10.6397 | -34.4837 | -18.0117
10.6458 | -34.4841 | -18.021
10.6519 | -34.4845 | -18.0304
10.658 | -34.4849 | -18.0397
10.6641 | -34.4853 | -18.0491
10.6702 | -34.4857 | -18.0584
10.6764 | -34.4861 | -18.0677
10.6825 | -34.4865 | -18.0769
10.6886 | -34.4869 | -18.0861
10.6947 | -34.4873 | -18.0954
10.7008 | -34.4878 | -18.1047
10.7069 | -34.4882 | -18.1141
10.713 | -34.4886 | -18.1234
10.7191 | -34.489 | -18.1328
10.7252 | -34.4894 | -18.1418
10.7313 | -34.4898 | -18.151
10.7374 | -34.4902 | -18.1604
10.7435 | -34.4906 | -18.1697
10.7496 | -34.491 | -18.1787
10.7557 | -34.4914 | -18.1881
10.7618 | -34.4918 | -18.1972
10.7679 | -34.4922 | -18.2065
10.774 | -34.4926 | -18.2157
10.7801 | -34.493 | -18.2251
10.7862 | -34.4934 | -18.2343
10.7923 | -34.4938 | -18.2436
10.7984 | -34.4942 | -18.2527
10.8045 | -34.4946 | -18.2618
10.8106 | -34.495 | -18.2713
10.8168 | -34.4954 | -18.2804
10.8229 | -34.4958 | -18.2896
10.829 | -34.4962 | -18.2988
10.8351 | -34.4966 | -18.3077
10.8412 | -34.497 | -18.3171
10.8473 | -34.4974 | -18.3261
10.8534 | -34.4978 | -18.3356
10.8595 | -34.4983 | -18.3447
10.8656 | -34.4987 | -18.3537
10.8717 | -34.499 | -18.3629
10.8778 | -34.4995 | -18.3722
10.8839 | -34.4999 | -18.3815
10.89 | -34.5003 | -18.3908
10.8961 | -34.5007 | -18.3999
10.9022 | -34.5011 | -18.409
10.9083 | -34.5015 | -18.418
10.9144 | -34.5019 | -18.4271
10.9205 | -34.5023 | -18.4364
10.9266 | -34.5027 | -18.4456
10.9327 | -34.5031 | -18.4548
10.9388 | -34.5036 | -18.4639
10.9449 | -34.504 | -18.4731
10.951 | -34.5044 | -18.4823
10.9571 | -34.5048 | -18.4914
10.9633 | -34.5052 | -18.5007
10.9694 | -34.5056 | -18.5102
10.9755 | -34.506 | -18.5193
10.9816 | -34.5064 | -18.5286
10.9877 | -34.5068 | -18.5376
10.9938 | -34.5072 | -18.547
10.9999 | -34.5076 | -18.556
11.006 | -34.5081 | -18.565
11.0121 | -34.5085 | -18.5742
11.0182 | -34.5089 | -18.5834
11.0243 | -34.5093 | -18.5925
11.0304 | -34.5097 | -18.6016
11.0365 | -34.5101 | -18.6106
11.0426 | -34.5105 | -18.6196
11.0487 | -34.511 | -18.6288
11.0548 | -34.5114 | -18.6378
11.0609 | -34.5118 | -18.647
11.067 | -34.5122 | -18.6562
11.0731 | -34.5126 | -18.6652
11.0792 | -34.5131 | -18.6742
11.0853 | -34.5135 | -18.6834
11.0914 | -34.5139 | -18.6927
11.0975 | -34.5143 | -18.7018
11.1037 | -34.5147 | -18.7109
11.1098 | -34.5151 | -18.7202
11.1159 | -34.5155 | -18.7292
11.122 | -34.5159 | -18.7382
11.1281 | -34.5163 | -18.7474
11.1342 | -34.5167 | -18.7565
11.1403 | -34.5172 | -18.7657
11.1464 | -34.5176 | -18.7748
11.1525 | -34.518 | -18.7837
11.1586 | -34.5184 | -18.793
11.1647 | -34.5188 | -18.802
11.1708 | -34.5192 | -18.8112
11.1769 | -34.5196 | -18.8203
11.183 | -34.5201 | -18.8293
11.1891 | -34.5205 | -18.8384
11.1952 | -34.5209 | -18.8475
11.2013 | -34.5213 | -18.8565
11.2074 | -34.5217 | -18.8656
11.2135 | -34.5221 | -18.8749
11.2196 | -34.5226 | -18.8839
11.2257 | -34.523 | -18.8931
11.2318 | -34.5234 | -18.9021
11.2379 | -34.5239 | -18.9111
11.244 | -34.5243 | -18.9203
11.2502 | -34.5247 | -18.9294
11.2563 | -34.5252 | -18.9384
11.2624 | -34.5256 | -18.9474
11.2685 | -34.526 | -18.9564
11.2746 | -34.5264 | -18.9654
11.2807 | -34.5268 | -18.9747
11.2868 | -34.5273 | -18.984
11.2929 | -34.5277 | -18.9928
11.299 | -34.5281 | -19.002
11.3051 | -34.5285 | -19.0112
11.3112 | -34.5289 | -19.0205
11.3173 | -34.5294 | -19.0295
11.3234 | -34.5298 | -19.0383
11.3295 | -34.5302 | -19.0476
11.3356 | -34.5306 | -19.0567
11.3417 | -34.531 | -19.0658
11.3478 | -34.5314 | -19.075
11.3539 | -34.5318 | -19.0841
11.36 | -34.5322 | -19.0931
11.3661 | -34.5327 | -19.102
11.3722 | -34.5331 | -19.111
11.3783 | -34.5335 | -19.12
11.3844 | -34.5339 | -19.129
11.3906 | -34.5344 | -19.138
11.3967 | -34.5348 | -19.147
11.4028 | -34.5352 | -19.1563
11.4089 | -34.5356 | -19.1651
11.415 | -34.5361 | -19.1742
11.4211 | -34.5365 | -19.1834
11.4272 | -34.5369 | -19.1924
11.4333 | -34.5373 | -19.2012
11.4394 | -34.5377 | -19.2103
11.4455 | -34.5381 | -19.2193
11.4516 | -34.5386 | -19.2284
11.4577 | -34.539 | -19.2374
11.4638 | -34.5394 | -19.2463
11.4699 | -34.5399 | -19.2554
11.476 | -34.5403 | -19.2646
11.4821 | -34.5407 | -19.2736
11.4882 | -34.5411 | -19.2827
11.4943 | -34.5415 | -19.2918
11.5004 | -34.542 | -19.3007
11.5065 | -34.5424 | -19.3097
11.5126 | -34.5428 | -19.3187
11.5187 | -34.5432 | -19.3278
11.5248 | -34.5436 | -19.337
11.5309 | -34.5441 | -19.3458
11.5371 | -34.5445 | -19.3549
11.5432 | -34.545 | -19.364
11.5493 | -34.5454 | -19.3731
11.5554 | -34.5458 | -19.3822
11.5615 | -34.5462 | -19.3912
11.5676 | -34.5466 | -19.4003
11.5737 | -34.5471 | -19.4091
11.5798 | -34.5475 | -19.418
11.5859 | -34.548 | -19.4271
11.592 | -34.5483 | -19.4363
11.5981 | -34.5488 | -19.4451
11.6042 | -34.5492 | -19.4539
11.6103 | -34.5496 | -19.463
11.6164 | -34.5501 | -19.4719
11.6225 | -34.5505 | -19.4808
11.6286 | -34.5509 | -19.4898
11.6347 | -34.5513 | -19.499
11.6408 | -34.5518 | -19.5078
11.6469 | -34.5522 | -19.5168
11.653 | -34.5526 | -19.5256
11.6591 | -34.553 | -19.5345
11.6652 | -34.5535 | -19.5435
11.6713 | -34.5539 | -19.5525
11.6775 | -34.5543 | -19.5612
11.6836 | -34.5547 | -19.5706
11.6897 | -34.5551 | -19.5794
11.6958 | -34.5555 | -19.5884
11.7019 | -34.556 | -19.5976
11.708 | -34.5564 | -19.6063
11.7141 | -34.5569 | -19.6152
11.7202 | -34.5573 | -19.6242
11.7263 | -34.5577 | -19.6332
11.7324 | -34.5581 | -19.6423
11.7385 | -34.5585 | -19.6511
11.7446 | -34.5589 | -19.6599
11.7507 | -34.5593 | -19.6686
11.7568 | -34.5597 | -19.6775
11.7629 | -34.5602 | -19.6865
11.769 | -34.5606 | -19.6954
11.7751 | -34.561 | -19.7042
11.7812 | -34.5615 | -19.713
11.7873 | -34.5619 | -19.7219
11.7934 | -34.5624 | -19.7308
11.7995 | -34.5628 | -19.7397
11.8056 | -34.5632 | -19.7486
11.8117 | -34.5636 | -19.7575
11.8178 | -34.5641 | -19.7665
11.824 | -34.5646 | -19.7753
11.8301 | -34.565 | -19.7843
11.8362 | -34.5654 | -19.7931
11.8423 | -34.5658 | -19.8019
11.8484 | -34.5662 | -19.8109
11.8545 | -34.5667 | -19.8196
11.8606 | -34.5671 | -19.8285
11.8667 | -34.5676 | -19.8375
11.8728 | -34.568 | -19.8462
11.8789 | -34.5684 | -19.8551
11.885 | -34.5689 | -19.8638
11.8911 | -34.5693 | -19.8726
11.8972 | -34.5697 | -19.8814
11.9033 | -34.5702 | -19.8903
11.9094 | -34.5706 | -19.8991
11.9155 | -34.571 | -19.9081
11.9216 | -34.5715 | -19.9168
11.9277 | -34.5719 | -19.9256
11.9338 | -34.5724 | -19.9345
11.9399 | -34.5728 | -19.9435
11.946 | -34.5733 | -19.9522
11.9521 | -34.5737 | -19.9611
11.9582 | -34.5741 | -19.9701
11.9644 | -34.5746 | -19.979
11.9705 | -34.575 | -19.9878
11.9766 | -34.5755 | -19.9966
11.9827 | -34.5759 | -20.0053
11.9888 | -34.5764 | -20.0143
11.9949 | -34.5768 | -20.0227
12.001 | -34.5772 | -20.0316
12.0071 | -34.5777 | -20.0402
12.0132 | -34.5781 | -20.049
12.0193 | -34.5785 | -20.0579
12.0254 | -34.5789 | -20.0668
12.0315 | -34.5794 | -20.0754
12.0376 | -34.5798 | -20.0841
12.0437 | -34.5802 | -20.0929
12.0498 | -34.5807 | -20.1018
12.0559 | -34.5811 | -20.1105
12.062 | -34.5816 | -20.119
12.0681 | -34.582 | -20.1278
12.0742 | -34.5824 | -20.1366
12.0803 | -34.5829 | -20.1454
12.0864 | -34.5833 | -20.1543
12.0925 | -34.5838 | -20.1631
12.0986 | -34.5843 | -20.1717
12.1047 | -34.5847 | -20.1806
12.1109 | -34.5851 | -20.1892
12.117 | -34.5856 | -20.198
12.1231 | -34.586 | -20.2068
12.1292 | -34.5864 | -20.2154
12.1353 | -34.5869 | -20.2243
12.1414 | -34.5873 | -20.2332
12.1475 | -34.5878 | -20.2422
12.1536 | -34.5882 | -20.251
12.1597 | -34.5886 | -20.2597
12.1658 | -34.5891 | -20.2683
12.1719 | -34.5896 | -20.2769
12.178 | -34.59 | -20.2856
12.1841 | -34.5905 | -20.2944
12.1902 | -34.5909 | -20.3031
12.1963 | -34.5914 | -20.3118
12.2024 | -34.5918 | -20.3205
12.2085 | -34.5923 | -20.3292
12.2146 | -34.5927 | -20.3379
12.2207 | -34.5931 | -20.3467
12.2268 | -34.5936 | -20.3555
12.2329 | -34.5941 | -20.3642
12.239 | -34.5946 | -20.3726
12.2451 | -34.595 | -20.3814
12.2513 | -34.5955 | -20.3902
12.2574 | -34.5959 | -20.399
12.2635 | -34.5963 | -20.408
12.2696 | -34.5968 | -20.4166
12.2757 | -34.5973 | -20.425
12.2818 | -34.5977 | -20.4336
12.2879 | -34.5982 | -20.4424
12.294 | -34.5986 | -20.4509
12.3001 | -34.5991 | -20.4596
12.3062 | -34.5996 | -20.4682
12.3123 | -34.6 | -20.4769
12.3184 | -34.6004 | -20.4854
12.3245 | -34.6008 | -20.4941
12.3306 | -34.6013 | -20.5027
12.3367 | -34.6017 | -20.5114
12.3428 | -34.6022 | -20.5202
12.3489 | -34.6027 | -20.5285
12.355 | -34.6031 | -20.5374
12.3611 | -34.6035 | -20.5461
12.3672 | -34.604 | -20.5547
12.3733 | -34.6044 | -20.5636
12.3794 | -34.6048 | -20.5722
12.3855 | -34.6053 | -20.5808
12.3916 | -34.6057 | -20.5895
12.3978 | -34.6062 | -20.5983
12.4039 | -34.6067 | -20.6069
12.41 | -34.6071 | -20.6155
12.4161 | -34.6075 | -20.6238
12.4222 | -34.608 | -20.6323
12.4283 | -34.6084 | -20.6409
12.4344 | -34.6089 | -20.6495
12.4405 | -34.6094 | -20.6585
12.4466 | -34.6098 | -20.667
12.4527 | -34.6103 | -20.6757
12.4588 | -34.6107 | -20.6845
12.4649 | -34.6112 | -20.6932
12.471 | -34.6116 | -20.7019
12.4771 | -34.6121 | -20.7105
12.4832 | -34.6125 | -20.7192
12.4893 | -34.613 | -20.7277
12.4954 | -34.6134 | -20.7363
12.5015 | -34.6139 | -20.745
12.5076 | -34.6144 | -20.7529
12.5137 | -34.615 | -20.7613
12.5198 | -34.6154 | -20.7692
12.5259 | -34.6157 | -20.778
12.532 | -34.6161 | -20.7856
12.5382 | -34.6166 | -20.7938
12.5443 | -34.617 | -20.8025
12.5504 | -34.6174 | -20.8112
12.5565 | -34.6179 | -20.8191
12.5626 | -34.6184 | -20.828
12.5687 | -34.6188 | -20.8366
12.5748 | -34.6193 | -20.8453
12.5809 | -34.6197 | -20.8539
12.587 | -34.6202 | -20.8625
12.5931 | -34.6207 | -20.871
12.5992 | -34.6212 | -20.8795
12.6053 | -34.6216 | -20.8878
12.6114 | -34.622 | -20.8962
12.6175 | -34.6225 | -20.9048
12.6236 | -34.6229 | -20.9136
12.6297 | -34.6234 | -20.9221
12.6358 | -34.6238 | -20.9306
12.6419 | -34.6243 | -20.9392
12.648 | -34.6248 | -20.9478
12.6541 | -34.6252 | -20.9563
12.6602 | -34.6256 | -20.9646
12.6663 | -34.6261 | -20.9732
12.6724 | -34.6265 | -20.9817
12.6785 | -34.627 | -20.9903
12.6847 | -34.6274 | -20.9985
12.6908 | -34.6279 | -21.0069
12.6969 | -34.6283 | -21.015
12.703 | -34.6287 | -21.0239
12.7091 | -34.6292 | -21.0327
12.7152 | -34.6296 | -21.0407
12.7213 | -34.6301 | -21.0492
12.7274 | -34.6306 | -21.0579
12.7335 | -34.631 | -21.0658
12.7396 | -34.6315 | -21.0746
12.7457 | -34.6319 | -21.0832
12.7518 | -34.6323 | -21.0915
12.7579 | -34.6328 | -21.1001
12.764 | -34.6332 | -21.1087
12.7701 | -34.6337 | -21.1174
12.7762 | -34.6341 | -21.126
12.7823 | -34.6346 | -21.1345
12.7884 | -34.6351 | -21.1431
12.7945 | -34.6355 | -21.1514
12.8006 | -34.6361 | -21.1596
12.8067 | -34.6365 | -21.1677
12.8128 | -34.6369 | -21.1761
12.8189 | -34.6374 | -21.1846
12.8251 | -34.6378 | -21.1932
12.8312 | -34.6382 | -21.2015
12.8373 | -34.6387 | -21.21
12.8434 | -34.6392 | -21.2186
12.8495 | -34.6397 | -21.227
12.8556 | -34.6401 | -21.2354
12.8617 | -34.6406 | -21.2441
12.8678 | -34.6411 | -21.2521
12.8739 | -34.6415 | -21.2609
12.88 | -34.6419 | -21.2684
12.8861 | -34.6423 | -21.2778
12.8922 | -34.6428 | -21.2866
12.8983 | -34.6434 | -21.295
12.9044 | -34.6439 | -21.3032
12.9105 | -34.6442 | -21.3122
12.9166 | -34.6447 | -21.321
12.9227 | -34.6452 | -21.3288
12.9288 | -34.6457 | -21.3372
12.9349 | -34.6462 | -21.3459
12.941 | -34.6466 | -21.3545
12.9471 | -34.647 | -21.3631
12.9532 | -34.6475 | -21.3712
12.9593 | -34.648 | -21.3795
12.9654 | -34.6484 | -21.3879
12.9716 | -34.6489 | -21.3965
12.9777 | -34.6494 | -21.4051
12.9838 | -34.6498 | -21.4136
12.9899 | -34.6502 | -21.4223
12.996 | -34.6507 | -21.4307
13.0021 | -34.6511 | -21.4392
13.0082 | -34.6516 | -21.4472
13.0143 | -34.652 | -21.4553
13.0204 | -34.6524 | -21.4637
13.0265 | -34.6529 | -21.4721
13.0326 | -34.6534 | -21.4808
13.0387 | -34.6538 | -21.4894
13.0448 | -34.6542 | -21.4977
13.0509 | -34.6547 | -21.5061
13.057 | -34.6551 | -21.5139
13.0631 | -34.6556 | -21.5223
13.0692 | -34.6561 | -21.5307
13.0753 | -34.6566 | -21.5394
13.0814 | -34.657 | -21.5478
13.0875 | -34.6575 | -21.556
13.0936 | -34.6579 | -21.5644
13.0997 | -34.6583 | -21.5728
13.1058 | -34.6588 | -21.581
13.112 | -34.6592 | -21.5893
13.1181 | -34.6597 | -21.598
13.1242 | -34.6601 | -21.6065
13.1303 | -34.6606 | -21.6149
13.1364 | -34.6609 | -21.6231
13.1425 | -34.6613 | -21.6315
13.1486 | -34.6618 | -21.64
13.1547 | -34.6623 | -21.649
13.1608 | -34.6627 | -21.6575
13.1669 | -34.6632 | -21.6658
13.173 | -34.6637 | -21.6743
13.1791 | -34.6642 | -21.6826
13.1852 | -34.6647 | -21.691
13.1913 | -34.6651 | -21.6997
13.1974 | -34.6656 | -21.7083
13.2035 | -34.6661 | -21.7166
13.2096 | -34.6666 | -21.7248
13.2157 | -34.6671 | -21.7333
13.2218 | -34.6675 | -21.7417
13.2279 | -34.668 | -21.7495
13.234 | -34.6684 | -21.7575
13.2401 | -34.6689 | -21.7657
13.2462 | -34.6693 | -21.7741
13.2524 | -34.6697 | -21.7828
13.2585 | -34.6702 | -21.7907
13.2646 | -34.6707 | -21.7993
13.2707 | -34.6712 | -21.8076
13.2768 | -34.6716 | -21.8159
13.2829 | -34.6721 | -21.8244
13.289 | -34.6726 | -21.8326
13.2951 | -34.673 | -21.8404
13.3012 | -34.6735 | -21.8483
13.3073 | -34.674 | -21.8565
13.3134 | -34.6744 | -21.8655
13.3195 | -34.6749 | -21.8734
13.3256 | -34.6754 | -21.8817
13.3317 | -34.6761 | -21.8889
13.3378 | -34.6767 | -21.8976
13.3439 | -34.6772 | -21.9059
13.35 | -34.6775 | -21.9151
13.3561 | -34.6783 | -21.9225
13.3622 | -34.6788 | -21.9331
13.3683 | -34.6795 | -21.9408
13.3744 | -34.68 | -21.9497
13.3805 | -34.6806 | -21.9586
13.3866 | -34.6809 | -21.9664
13.3927 | -34.6813 | -21.9745
13.3989 | -34.682 | -21.9839
13.405 | -34.6823 | -21.9918
13.4111 | -34.6829 | -22.0007
13.4172 | -34.6833 | -22.0089
13.4233 | -34.6837 | -22.0172
13.4294 | -34.6841 | -22.0263
13.4355 | -34.6848 | -22.0348
13.4416 | -34.6853 | -22.0431
13.4477 | -34.6857 | -22.0516
13.4538 | -34.6863 | -22.0598
13.4599 | -34.6869 | -22.0677
13.466 | -34.6872 | -22.0757
13.4721 | -34.6877 | -22.0848
13.4782 | -34.6882 | -22.0925
13.4843 | -34.6886 | -22.1008
13.4904 | -34.6892 | -22.1092
13.4965 | -34.6898 | -22.1179
13.5026 | -34.6902 | -22.1266
13.5087 | -34.6907 | -22.1355
13.5148 | -34.6913 | -22.1434
13.5209 | -34.6917 | -22.1523
13.527 | -34.6921 | -22.1608
13.5331 | -34.6927 | -22.1683
13.5393 | -34.6932 | -22.1757
13.5454 | -34.6937 | -22.184
13.5515 | -34.6942 | -22.1927
13.5576 | -34.6946 | -22.2003
13.5637 | -34.6951 | -22.2079
13.5698 | -34.6957 | -22.2159
13.5759 | -34.6962 | -22.2245
13.582 | -34.6966 | -22.2325
13.5881 | -34.6971 | -22.2412
13.5942 | -34.6976 | -22.2497
13.6003 | -34.6982 | -22.2583
13.6064 | -34.6986 | -22.2672
13.6125 | -34.6991 | -22.2759
13.6186 | -34.6995 | -22.2839
13.6247 | -34.7 | -22.2924
13.6308 | -34.7005 | -22.3007
13.6369 | -34.7009 | -22.3094
13.643 | -34.7014 | -22.3182
13.6491 | -34.702 | -22.3268
13.6552 | -34.7025 | -22.3349
13.6613 | -34.703 | -22.3426
13.6674 | -34.7035 | -22.3501
13.6735 | -34.7039 | -22.3593
13.6796 | -34.7044 | -22.3668
13.6858 | -34.7049 | -22.3751
13.6919 | -34.7054 | -22.3833
13.698 | -34.7059 | -22.3922
13.7041 | -34.7063 | -22.4005
13.7102 | -34.7068 | -22.4089
13.7163 | -34.7074 | -22.4179
13.7224 | -34.708 | -22.4265
13.7285 | -34.7085 | -22.4347
13.7346 | -34.709 | -22.4434
13.7407 | -34.7095 | -22.4518
13.7468 | -34.71 | -22.4595
13.7529 | -34.7106 | -22.4678
13.759 | -34.7111 | -22.4762
13.7651 | -34.7116 | -22.4845
13.7712 | -34.7121 | -22.4928
13.7773 | -34.7125 | -22.501
13.7834 | -34.7129 | -22.5086
13.7895 | -34.7136 | -22.5171
13.7956 | -34.7139 | -22.5254
13.8017 | -34.7143 | -22.5344
13.8078 | -34.7149 | -22.5427
13.8139 | -34.7154 | -22.5503
13.82 | -34.7158 | -22.5596
13.8262 | -34.7162 | -22.5678
13.8323 | -34.7166 | -22.576
13.8384 | -34.7171 | -22.584
13.8445 | -34.7177 | -22.5916
13.8506 | -34.7183 | -22.6005
13.8567 | -34.719 | -22.6092
13.8628 | -34.7195 | -22.6174
13.8689 | -34.7198 | -22.6249
13.875 | -34.7204 | -22.6328
13.8811 | -34.7207 | -22.6404
13.8872 | -34.7213 | -22.6494
13.8933 | -34.722 | -22.6574
13.8994 | -34.7224 | -22.6657
13.9055 | -34.7229 | -22.6736
13.9116 | -34.7234 | -22.6821
13.9177 | -34.7239 | -22.6911
13.9238 | -34.7243 | -22.6993
13.9299 | -34.7249 | -22.707
13.936 | -34.7253 | -22.7156
13.9421 | -34.7259 | -22.7236
13.9482 | -34.7263 | -22.732
13.9543 | -34.7267 | -22.7402
13.9604 | -34.7272 | -22.7488
13.9665 | -34.7278 | -22.7565
13.9727 | -34.7283 | -22.7653
13.9788 | -34.7289 | -22.7741
13.9849 | -34.7293 | -22.7822
13.991 | -34.7297 | -22.7899
13.9971 | -34.7304 | -22.798
14.0032 | -34.7308 | -22.8067
14.0093 | -34.7314 | -22.8145
14.0154 | -34.732 | -22.823
14.0215 | -34.7325 | -22.8305
14.0276 | -34.7332 | -22.8389
14.0337 | -34.7338 | -22.8472
14.0398 | -34.7342 | -22.8556
14.0459 | -34.7346 | -22.8627
14.052 | -34.7352 | -22.8705
14.0581 | -34.7357 | -22.8794
14.0642 | -34.7362 | -22.8875
14.0703 | -34.7365 | -22.896
14.0764 | -34.737 | -22.903
14.0825 | -34.7375 | -22.911
14.0886 | -34.738 | -22.9195
14.0947 | -34.7384 | -22.9273
14.1008 | -34.7389 | -22.9358
14.1069 | -34.7392 | -22.9444
14.1131 | -34.7398 | -22.9529
14.1192 | -34.7403 | -22.9606
14.1253 | -34.7409 | -22.9691
14.1314 | -34.7413 | -22.9764
14.1375 | -34.742 | -22.9851
14.1436 | -34.7424 | -22.9933
14.1497 | -34.7428 | -23.0016
14.1558 | -34.7433 | -23.0103
14.1619 | -34.7439 | -23.019
14.168 | -34.7444 | -23.0261
14.1741 | -34.7447 | -23.03
14.1802 | -34.7454 | -23.0355
14.1863 | -34.7457 | -23.0442
14.1924 | -34.7461 | -23.0536
14.1985 | -34.7466 | -23.0622
14.2046 | -34.7471 | -23.0683
14.2107 | -34.7475 | -23.0735
14.2168 | -34.7481 | -23.0812
14.2229 | -34.7485 | -23.0894
14.229 | -34.749 | -23.0968
14.2351 | -34.7496 | -23.104
14.2412 | -34.7502 | -23.1122
14.2473 | -34.7508 | -23.1203
14.2534 | -34.7513 | -23.1283
14.2596 | -34.7518 | -23.1359
14.2657 | -34.7523 | -23.1439
14.2718 | -34.753 | -23.1527
14.2779 | -34.7535 | -23.1608
14.284 | -34.7541 | -23.1696
14.2901 | -34.7547 | -23.1779
14.2962 | -34.7552 | -23.1858
14.3023 | -34.7557 | -23.1936
14.3084 | -34.7562 | -23.202
14.3145 | -34.7567 | -23.2098
14.3206 | -34.7572 | -23.2182
14.3267 | -34.7578 | -23.2263
14.3328 | -34.7583 | -23.2342
14.3389 | -34.7588 | -23.2426
14.345 | -34.7593 | -23.2505
14.3511 | -34.7598 | -23.2586
14.3572 | -34.7602 | -23.2666
14.3633 | -34.7607 | -23.2744
14.3694 | -34.7613 | -23.2825
14.3755 | -34.7617 | -23.2906
14.3816 | -34.7622 | -23.2988
14.3877 | -34.7627 | -23.3065
14.3938 | -34.7632 | -23.3147
14.4 | -34.7637 | -23.3227
14.4061 | -34.7642 | -23.3307
14.4122 | -34.7647 | -23.3386
14.4183 | -34.7652 | -23.3464
14.4244 | -34.7657 | -23.3538
14.4305 | -34.7663 | -23.3618
14.4366 | -34.7668 | -23.3697
14.4427 | -34.7673 | -23.3778
14.4488 | -34.7678 | -23.3859
14.4549 | -34.7684 | -23.3935
14.461 | -34.7689 | -23.4012
14.4671 | -34.7693 | -23.4093
14.4732 | -34.7698 | -23.4173
14.4793 | -34.7703 | -23.4256
14.4854 | -34.7708 | -23.4334
14.4915 | -34.7713 | -23.4404
14.4976 | -34.7718 | -23.4482
14.5037 | -34.7723 | -23.4557
14.5098 | -34.7727 | -23.4634
14.5159 | -34.7732 | -23.4714
14.522 | -34.7737 | -23.4795
14.5281 | -34.7741 | -23.4869
14.5342 | -34.7747 | -23.4948
14.5403 | -34.7752 | -23.5027
14.5465 | -34.7757 | -23.5107
14.5526 | -34.7762 | -23.5189
14.5587 | -34.7767 | -23.5269
14.5648 | -34.7773 | -23.5344
14.5709 | -34.7778 | -23.5423
14.577 | -34.7783 | -23.5499
14.5831 | -34.7788 | -23.5578
14.5892 | -34.7793 | -23.5661
14.5953 | -34.7799 | -23.5738
14.6014 | -34.7804 | -23.5814
14.6075 | -34.7809 | -23.5893
14.6136 | -34.7814 | -23.5973
14.6197 | -34.7819 | -23.6051
14.6258 | -34.7825 | -23.613
14.6319 | -34.783 | -23.6207
14.638 | -34.7835 | -23.6284
14.6441 | -34.784 | -23.6364
14.6502 | -34.7845 | -23.6443
14.6563 | -34.785 | -23.6522
14.6624 | -34.7855 | -23.6599
14.6685 | -34.786 | -23.668
14.6746 | -34.7866 | -23.6756
14.6807 | -34.7871 | -23.6836
14.6869 | -34.7876 | -23.6914
14.693 | -34.7881 | -23.6989
14.6991 | -34.7886 | -23.7065
14.7052 | -34.7892 | -23.7135
14.7113 | -34.7897 | -23.7215
14.7174 | -34.7902 | -23.7294
14.7235 | -34.7907 | -23.7372
14.7296 | -34.7912 | -23.7448
14.7357 | -34.7917 | -23.752
14.7418 | -34.7921 | -23.7603
14.7479 | -34.7926 | -23.7677
14.754 | -34.7932 | -23.7756
14.7601 | -34.7937 | -23.7834
14.7662 | -34.7942 | -23.7907
14.7723 | -34.7947 | -23.7983
14.7784 | -34.7952 | -23.8063
14.7845 | -34.7958 | -23.8144
14.7906 | -34.7962 | -23.8224
14.7967 | -34.7967 | -23.8301
14.8028 | -34.7972 | -23.8377
14.8089 | -34.7977 | -23.8452
14.815 | -34.7982 | -23.8528
14.8211 | -34.7987 | -23.8605
14.8272 | -34.7992 | -23.8687
14.8334 | -34.7997 | -23.8767
14.8395 | -34.8002 | -23.8841
14.8456 | -34.8007 | -23.8914
14.8517 | -34.8012 | -23.8992
14.8578 | -34.8017 | -23.9066
14.8639 | -34.8022 | -23.9143
14.87 | -34.8027 | -23.9217
14.8761 | -34.8032 | -23.9291
14.8822 | -34.8036 | -23.9371
14.8883 | -34.8041 | -23.9447
14.8944 | -34.8046 | -23.9522
14.9005 | -34.8051 | -23.9601
14.9066 | -34.8057 | -23.9675
14.9127 | -34.8061 | -23.9743
14.9188 | -34.8066 | -23.982
14.9249 | -34.8071 | -23.9897
14.931 | -34.8076 | -23.9979
14.9371 | -34.8081 | -24.005
14.9432 | -34.8087 | -24.0123
14.9493 | -34.8091 | -24.0199
14.9554 | -34.8096 | -24.0279
14.9615 | -34.81 | -24.0356
14.9676 | -34.8105 | -24.0433
14.9738 | -34.8111 | -24.0505
14.9799 | -34.8115 | -24.0586
14.986 | -34.812 | -24.0657
14.9921 | -34.8125 | -24.0731
14.9982 | -34.813 | -24.0806
15.0043 | -34.8136 | -24.0878
15.0104 | -34.8139 | -24.0956
15.0165 | -34.8145 | -24.1028
15.0226 | -34.815 | -24.1114
15.0287 | -34.8153 | -24.1203
15.0348 | -34.8156 | -24.1291
15.0409 | -34.8161 | -24.1364
15.047 | -34.8167 | -24.1445
15.0531 | -34.8171 | -24.1519
15.0592 | -34.8176 | -24.1594
15.0653 | -34.8181 | -24.1672
15.0714 | -34.8186 | -24.1746
15.0775 | -34.8191 | -24.1823
15.0836 | -34.8196 | -24.1898
15.0897 | -34.8201 | -24.1969
15.0958 | -34.8206 | -24.2043
15.1019 | -34.8211 | -24.2118
15.108 | -34.8216 | -24.2191
15.1141 | -34.822 | -24.2265
15.1203 | -34.8225 | -24.2344
15.1264 | -34.823 | -24.2421
15.1325 | -34.8235 | -24.2497
15.1386 | -34.824 | -24.2573
15.1447 | -34.8245 | -24.2649
15.1508 | -34.825 | -24.2729
15.1569 | -34.8255 | -24.2806
15.163 | -34.826 | -24.2878
15.1691 | -34.8265 | -24.2952
15.1752 | -34.827 | -24.3026
15.1813 | -34.8274 | -24.3102
15.1874 | -34.8279 | -24.318
15.1935 | -34.8283 | -24.3248
15.1996 | -34.8288 | -24.3323
15.2057 | -34.8293 | -24.3395
15.2118 | -34.8298 | -24.3473
15.2179 | -34.8302 | -24.3549
15.224 | -34.8307 | -24.3629
15.2301 | -34.8312 | -24.3699
15.2362 | -34.8317 | -24.3774
15.2423 | -34.8321 | -24.3848
15.2484 | -34.8326 | -24.3918
15.2545 | -34.8331 | -24.3995
15.2607 | -34.8336 | -24.4071
15.2668 | -34.8339 | -24.4143
15.2729 | -34.8345 | -24.4218
15.279 | -34.8349 | -24.4293
15.2851 | -34.8353 | -24.4368
15.2912 | -34.8358 | -24.4444
15.2973 | -34.8363 | -24.4519
15.3034 | -34.8367 | -24.4589
15.3095 | -34.8372 | -24.4659
15.3156 | -34.8377 | -24.4735
15.3217 | -34.8382 | -24.4809
15.3278 | -34.8387 | -24.4883
15.3339 | -34.8391 | -24.4955
15.34 | -34.8396 | -24.5028
15.3461 | -34.8401 | -24.5103
15.3522 | -34.8406 | -24.5178
15.3583 | -34.841 | -24.5254
15.3644 | -34.8415 | -24.5328
15.3705 | -34.842 | -24.54
15.3766 | -34.8425 | -24.5475
15.3827 | -34.843 | -24.5548
15.3888 | -34.8435 | -24.5622
15.3949 | -34.844 | -24.5697
15.401 | -34.8445 | -24.5773
15.4072 | -34.8449 | -24.5844
15.4133 | -34.8454 | -24.5918
15.4194 | -34.8458 | -24.5995
15.4255 | -34.8463 | -24.6066
15.4316 | -34.8467 | -24.6139
15.4377 | -34.8472 | -24.6214
15.4438 | -34.8476 | -24.6291
15.4499 | -34.848 | -24.6357
15.456 | -34.8485 | -24.6436
15.4621 | -34.849 | -24.6512
15.4682 | -34.8495 | -24.658
15.4743 | -34.8497 | -24.6653
15.4804 | -34.8501 | -24.6724
15.4865 | -34.8506 | -24.6801
15.4926 | -34.8509 | -24.6877
15.4987 | -34.8515 | -24.695
15.5048 | -34.8519 | -24.7023
15.5109 | -34.8524 | -24.7087
15.517 | -34.8528 | -24.7161
15.5231 | -34.8533 | -24.7235
15.5292 | -34.8537 | -24.7305
15.5353 | -34.8542 | -24.7376
15.5414 | -34.8547 | -24.7448
15.5476 | -34.8551 | -24.7514
15.5537 | -34.8556 | -24.7591
15.5598 | -34.856 | -24.7667
15.5659 | -34.8565 | -24.7743
15.572 | -34.8569 | -24.7819
15.5781 | -34.8574 | -24.7888
15.5842 | -34.8578 | -24.796
15.5903 | -34.8583 | -24.803
15.5964 | -34.8587 | -24.8107
15.6025 | -34.8592 | -24.8179
15.6086 | -34.8596 | -24.8253
15.6147 | -34.86 | -24.8324
15.6208 | -34.8605 | -24.8394
15.6269 | -34.8609 | -24.8468
15.633 | -34.8613 | -24.8543
15.6391 | -34.8618 | -24.8615
15.6452 | -34.8622 | -24.8688
15.6513 | -34.8627 | -24.8755
15.6574 | -34.8631 | -24.8831
15.6635 | -34.8636 | -24.8905
15.6696 | -34.864 | -24.8977
15.6757 | -34.8645 | -24.9053
15.6818 | -34.8649 | -24.9125
15.688 | -34.8653 | -24.9197
15.6941 | -34.8657 | -24.9269
15.7002 | -34.8662 | -24.9342
15.7063 | -34.8666 | -24.9413
15.7124 | -34.8671 | -24.9489
15.7185 | -34.8675 | -24.956
15.7246 | -34.8679 | -24.9638
15.7307 | -34.8684 | -24.9713
15.7368 | -34.8688 | -24.9786
15.7429 | -34.8693 | -24.9858
15.749 | -34.8698 | -24.9928
15.7551 | -34.8702 | -25.0004
15.7612 | -34.8706 | -25.0075
15.7673 | -34.871 | -25.0149
15.7734 | -34.8715 | -25.0221
15.7795 | -34.872 | -25.0294
15.7856 | -34.8725 | -25.0363
15.7917 | -34.8729 | -25.0435
15.7978 | -34.8733 | -25.0511
15.8039 | -34.8738 | -25.0586
15.81 | -34.8743 | -25.0658
15.8161 | -34.8747 | -25.0731
15.8222 | -34.8751 | -25.0807
15.8283 | -34.8757 | -25.0882
15.8345 | -34.8761 | -25.0953
15.8406 | -34.8766 | -25.103
15.8467 | -34.8769 | -25.1102
15.8528 | -34.8774 | -25.1165
15.8589 | -34.878 | -25.1232
15.865 | -34.8786 | -25.1311
15.8711 | -34.8789 | -25.1377
15.8772 | -34.8795 | -25.1447
15.8833 | -34.8801 | -25.1516
15.8894 | -34.8805 | -25.1592
15.8955 | -34.8809 | -25.1668
15.9016 | -34.8813 | -25.1741
15.9077 | -34.8818 | -25.182
15.9138 | -34.8822 | -25.1894
15.9199 | -34.8827 | -25.1969
15.926 | -34.8831 | -25.2042
15.9321 | -34.8835 | -25.2121
15.9382 | -34.884 | -25.2192
15.9443 | -34.8845 | -25.2267
15.9504 | -34.885 | -25.2342
15.9565 | -34.8854 | -25.2415
15.9626 | -34.8858 | -25.2491
15.9687 | -34.8863 | -25.2563
15.9749 | -34.8867 | -25.2635
15.981 | -34.8871 | -25.2704
15.9871 | -34.8875 | -25.278
15.9932 | -34.888 | -25.2853
15.9993 | -34.8884 | -25.293
16.0054 | -34.8888 | -25.2999
16.0115 | -34.8893 | -25.3072
16.0176 | -34.8897 | -25.3147
16.0237 | -34.8902 | -25.322
16.0298 | -34.8906 | -25.3295
16.0359 | -34.891 | -25.3367
16.042 | -34.8914 | -25.3439
16.0481 | -34.8919 | -25.3511
16.0542 | -34.8924 | -25.3586
16.0603 | -34.8928 | -25.3663
16.0664 | -34.8932 | -25.3733
16.0725 | -34.8936 | -25.3804
16.0786 | -34.894 | -25.3877
16.0847 | -34.8945 | -25.3951
16.0908 | -34.8949 | -25.4028
16.0969 | -34.8954 | -25.4102
16.103 | -34.8958 | -25.4175
16.1091 | -34.8963 | -25.4249
16.1152 | -34.8968 | -25.4322
16.1214 | -34.8973 | -25.4398
16.1275 | -34.8977 | -25.4473
16.1336 | -34.8981 | -25.4546
16.1397 | -34.8986 | -25.4621
16.1458 | -34.8989 | -25.4701
16.1519 | -34.8994 | -25.4776
16.158 | -34.8999 | -25.4855
16.1641 | -34.9003 | -25.4927
16.1702 | -34.9007 | -25.4998
16.1763 | -34.9012 | -25.5073
16.1824 | -34.9016 | -25.5147
16.1885 | -34.9021 | -25.522
16.1946 | -34.9026 | -25.5293
16.2007 | -34.903 | -25.5369
16.2068 | -34.9034 | -25.5441
16.2129 | -34.9038 | -25.5518
16.219 | -34.9043 | -25.5594
16.2251 | -34.9048 | -25.5668
16.2312 | -34.9052 | -25.5742
16.2373 | -34.9056 | -25.5814
16.2434 | -34.9061 | -25.5886
16.2495 | -34.9065 | -25.5956
16.2556 | -34.9069 | -25.6031
16.2618 | -34.9074 | -25.6102
16.2679 | -34.9078 | -25.6179
16.274 | -34.9083 | -25.6255
16.2801 | -34.9087 | -25.633
16.2862 | -34.9091 | -25.641
16.2923 | -34.9096 | -25.6485
16.2984 | -34.91 | -25.656
16.3045 | -34.9104 | -25.6633
16.3106 | -34.9108 | -25.6707
16.3167 | -34.9113 | -25.6782
16.3228 | -34.9117 | -25.6859
16.3289 | -34.9122 | -25.6935
16.335 | -34.9126 | -25.7006
16.3411 | -34.913 | -25.7081
16.3472 | -34.9134 | -25.7156
16.3533 | -34.9139 | -25.7229
16.3594 | -34.9144 | -25.7303
16.3655 | -34.9148 | -25.7377
16.3716 | -34.9152 | -25.7452
16.3777 | -34.9157 | -25.7528
16.3838 | -34.9161 | -25.7606
16.3899 | -34.9165 | -25.7682
16.396 | -34.9169 | -25.7754
16.4021 | -34.9174 | -25.7829
16.4083 | -34.9178 | -25.7904
16.4144 | -34.9182 | -25.7977
16.4205 | -34.9187 | -25.8053
16.4266 | -34.9192 | -25.813
16.4327 | -34.9196 | -25.8208
16.4388 | -34.92 | -25.8283
16.4449 | -34.9204 | -25.8356
16.451 | -34.9209 | -25.8434
16.4571 | -34.9213 | -25.8508
16.4632 | -34.9218 | -25.8586
16.4693 | -34.9222 | -25.8659
16.4754 | -34.9227 | -25.8736
16.4815 | -34.9231 | -25.8812
16.4876 | -34.9236 | -25.8884
16.4937 | -34.924 | -25.8961
16.4998 | -34.9245 | -25.9037
16.5059 | -34.9249 | -25.9112
16.512 | -34.9253 | -25.9187
16.5181 | -34.9258 | -25.9265
16.5242 | -34.9263 | -25.9339
16.5303 | -34.9268 | -25.9418
16.5364 | -34.9272 | -25.9493
16.5425 | -34.9276 | -25.9567
16.5487 | -34.928 | -25.9647
16.5548 | -34.9285 | -25.9722
16.5609 | -34.9289 | -25.9798
16.567 | -34.9294 | -25.9872
16.5731 | -34.9299 | -25.9952
16.5792 | -34.9302 | -26.0025
16.5853 | -34.9307 | -26.0101
16.5914 | -34.9312 | -26.0177
16.5975 | -34.9317 | -26.0253
16.6036 | -34.9321 | -26.0331
16.6097 | -34.9327 | -26.0402
16.6158 | -34.933 | -26.0479
16.6219 | -34.9335 | -26.0558
16.628 | -34.934 | -26.0633
16.6341 | -34.9344 | -26.0711
16.6402 | -34.9349 | -26.0789
16.6463 | -34.9353 | -26.0863
16.6524 | -34.9357 | -26.094
16.6585 | -34.9362 | -26.1018
16.6646 | -34.9367 | -26.1095
16.6707 | -34.9372 | -26.1173
16.6768 | -34.9377 | -26.1248
16.6829 | -34.9383 | -26.1322
16.689 | -34.9387 | -26.1398
16.6952 | -34.9391 | -26.1473
16.7013 | -34.9396 | -26.1553
16.7074 | -34.94 | -26.1626
16.7135 | -34.9405 | -26.1703
16.7196 | -34.941 | -26.1777
16.7257 | -34.9414 | -26.1853
16.7318 | -34.9419 | -26.1929
16.7379 | -34.9423 | -26.2004
16.744 | -34.9428 | -26.208
16.7501 | -34.9433 | -26.2158
16.7562 | -34.9437 | -26.2238
16.7623 | -34.9442 | -26.231
16.7684 | -34.9447 | -26.2386
16.7745 | -34.9451 | -26.2461
16.7806 | -34.9456 | -26.2537
16.7867 | -34.9461 | -26.2614
16.7928 | -34.9465 | -26.269
16.7989 | -34.9469 | -26.2766
16.805 | -34.9474 | -26.2843
16.8111 | -34.9479 | -26.292
16.8172 | -34.9484 | -26.2998
16.8233 | -34.9488 | -26.3074
16.8294 | -34.9493 | -26.3151
16.8356 | -34.9498 | -26.3229
16.8417 | -34.9502 | -26.3305
16.8478 | -34.9507 | -26.3383
16.8539 | -34.9512 | -26.3458
16.86 | -34.9516 | -26.3535
16.8661 | -34.9521 | -26.361
16.8722 | -34.9525 | -26.3688
16.8783 | -34.953 | -26.3765
16.8844 | -34.9535 | -26.3844
16.8905 | -34.954 | -26.3922
16.8966 | -34.9545 | -26.3999
16.9027 | -34.9549 | -26.4077
16.9088 | -34.9554 | -26.4153
16.9149 | -34.9559 | -26.423
16.921 | -34.9564 | -26.4308
16.9271 | -34.9568 | -26.4385
16.9332 | -34.9573 | -26.4461
16.9393 | -34.9577 | -26.4539
16.9454 | -34.9582 | -26.4616
16.9515 | -34.9587 | -26.4694
16.9576 | -34.9592 | -26.4773
16.9637 | -34.9596 | -26.4848
16.9698 | -34.9601 | -26.4924
16.9759 | -34.9605 | -26.5002
16.9821 | -34.961 | -26.508
16.9882 | -34.9615 | -26.5158
16.9943 | -34.962 | -26.5234
17.0004 | -34.9625 | -26.5311
17.0065 | -34.9629 | -26.5389
17.0126 | -34.9634 | -26.5464
17.0187 | -34.9638 | -26.5541
17.0248 | -34.9644 | -26.562
17.0309 | -34.9648 | -26.5701
17.037 | -34.9653 | -26.5775
17.0431 | -34.9658 | -26.5853
17.0492 | -34.9663 | -26.5931
17.0553 | -34.9668 | -26.6009
17.0614 | -34.9673 | -26.6088
17.0675 | -34.9677 | -26.6169
17.0736 | -34.9682 | -26.6248
17.0797 | -34.9687 | -26.6324
17.0858 | -34.9691 | -26.6401
17.0919 | -34.9697 | -26.6479
17.098 | -34.9702 | -26.6558
17.1041 | -34.9707 | -26.6634
17.1102 | -34.9712 | -26.6712
17.1163 | -34.9716 | -26.6789
17.1225 | -34.9721 | -26.6865
17.1286 | -34.9726 | -26.6946
17.1347 | -34.9731 | -26.7025
17.1408 | -34.9735 | -26.7104
17.1469 | -34.9741 | -26.7182
17.153 | -34.9745 | -26.7257
17.1591 | -34.975 | -26.7336
17.1652 | -34.9755 | -26.7414
17.1713 | -34.976 | -26.7493
17.1774 | -34.9765 | -26.7572
17.1835 | -34.977 | -26.7648
17.1896 | -34.9774 | -26.7723
17.1957 | -34.9779 | -26.7804
17.2018 | -34.9784 | -26.788
17.2079 | -34.9789 | -26.7958
17.214 | -34.9794 | -26.8036
17.2201 | -34.9799 | -26.8114
17.2262 | -34.9804 | -26.8192
17.2323 | -34.9809 | -26.827
17.2384 | -34.9814 | -26.8349
17.2445 | -34.9819 | -26.8424
17.2506 | -34.9824 | -26.8499
17.2567 | -34.9829 | -26.8578
17.2628 | -34.9834 | -26.8655
17.269 | -34.9839 | -26.8734
17.2751 | -34.9844 | -26.8811
17.2812 | -34.9849 | -26.8889
17.2873 | -34.9854 | -26.8969
17.2934 | -34.9859 | -26.9046
17.2995 | -34.9863 | -26.9125
17.3056 | -34.9869 | -26.9204
17.3117 | -34.9874 | -26.9285
17.3178 | -34.9878 | -26.9363
17.3239 | -34.9884 | -26.9442
17.33 | -34.9888 | -26.9521
17.3361 | -34.9893 | -26.9596
17.3422 | -34.9899 | -26.9673
17.3483 | -34.9904 | -26.9752
17.3544 | -34.9909 | -26.9829
17.3605 | -34.9913 | -26.9906
17.3666 | -34.9919 | -26.9983
17.3727 | -34.9923 | -27.0059
17.3788 | -34.9928 | -27.0137
17.3849 | -34.9933 | -27.0214
17.391 | -34.9938 | -27.0292
17.3971 | -34.9943 | -27.0369
17.4032 | -34.9948 | -27.0446
17.4094 | -34.9953 | -27.0523
17.4155 | -34.9958 | -27.0601
17.4216 | -34.9963 | -27.0677
17.4277 | -34.9968 | -27.0754
17.4338 | -34.9973 | -27.0832
17.4399 | -34.9978 | -27.0912
17.446 | -34.9983 | -27.099
17.4521 | -34.9988 | -27.1066
17.4582 | -34.9993 | -27.1143
17.4643 | -34.9998 | -27.122
17.4704 | -35.0003 | -27.13
17.4765 | -35.0008 | -27.1377
17.4826 | -35.0013 | -27.1456
17.4887 | -35.0018 | -27.1534
17.4948 | -35.0023 | -27.161
17.5009 | -35.0028 | -27.1689
17.507 | -35.0033 | -27.1769
17.5131 | -35.0039 | -27.1846
17.5192 | -35.0044 | -27.1924
17.5253 | -35.0049 | -27.1999
17.5314 | -35.0054 | -27.2078
17.5375 | -35.0059 | -27.2154
17.5436 | -35.0064 | -27.2234
17.5497 | -35.0069 | -27.2312
17.5559 | -35.0074 | -27.239
17.562 | -35.0079 | -27.2466
17.5681 | -35.0085 | -27.2546
17.5742 | -35.009 | -27.2624
17.5803 | -35.0095 | -27.2701
17.5864 | -35.0101 | -27.278
17.5925 | -35.0106 | -27.2859
17.5986 | -35.0111 | -27.2937
17.6047 | -35.0116 | -27.3017
17.6108 | -35.0121 | -27.3094
17.6169 | -35.0126 | -27.317
17.623 | -35.0132 | -27.3247
17.6291 | -35.0136 | -27.3326
17.6352 | -35.0142 | -27.3404
17.6413 | -35.0147 | -27.3485
17.6474 | -35.0153 | -27.3562
17.6535 | -35.0158 | -27.364
17.6596 | -35.0163 | -27.3719
17.6657 | -35.0168 | -27.3797
17.6718 | -35.0173 | -27.3877
17.6779 | -35.0178 | -27.3953
17.684 | -35.0183 | -27.4032
17.6901 | -35.0188 | -27.411
17.6963 | -35.0194 | -27.4188
17.7024 | -35.0199 | -27.4266
17.7085 | -35.0204 | -27.4343
17.7146 | -35.0209 | -27.4419
17.7207 | -35.0215 | -27.4497
17.7268 | -35.022 | -27.4573
17.7329 | -35.0225 | -27.4651
17.739 | -35.023 | -27.473
17.7451 | -35.0235 | -27.4807
17.7512 | -35.0241 | -27.4886
17.7573 | -35.0246 | -27.4965
17.7634 | -35.0251 | -27.5043
17.7695 | -35.0256 | -27.512
17.7756 | -35.0261 | -27.5203
17.7817 | -35.0265 | -27.5282
17.7878 | -35.0271 | -27.5362
17.7939 | -35.0276 | -27.5444
17.8 | -35.0281 | -27.5522
17.8061 | -35.0288 | -27.5601
17.8122 | -35.0293 | -27.5685
17.8183 | -35.0298 | -27.5767
17.8244 | -35.0303 | -27.5845
17.8305 | -35.0309 | -27.5923
17.8366 | -35.0314 | -27.5999
17.8428 | -35.0319 | -27.6081
17.8489 | -35.0325 | -27.6159
17.855 | -35.0329 | -27.6233
17.8611 | -35.0334 | -27.6311
17.8672 | -35.034 | -27.6388
17.8733 | -35.0345 | -27.6466
17.8794 | -35.035 | -27.6544
17.8855 | -35.0356 | -27.6621
17.8916 | -35.0361 | -27.6698
17.8977 | -35.0366 | -27.6775
17.9038 | -35.0371 | -27.685
17.9099 | -35.0376 | -27.6928
17.916 | -35.0381 | -27.7008
17.9221 | -35.0387 | -27.7086
17.9282 | -35.0392 | -27.7165
17.9343 | -35.0397 | -27.7241
17.9404 | -35.0402 | -27.7317
17.9465 | -35.0408 | -27.7397
17.9526 | -35.0413 | -27.7475
17.9587 | -35.0418 | -27.7554
17.9648 | -35.0423 | -27.7629
17.9709 | -35.0428 | -27.7706
17.977 | -35.0433 | -27.7784
17.9832 | -35.0438 | -27.7861
17.9893 | -35.0444 | -27.794
17.9954 | -35.0449 | -27.8017
18.0015 | -35.0454 | -27.8096
18.0076 | -35.046 | -27.8173
18.0137 | -35.0465 | -27.8251
18.0198 | -35.047 | -27.8329
18.0259 | -35.0475 | -27.8408
18.032 | -35.048 | -27.8487
18.0381 | -35.0486 | -27.8565
18.0442 | -35.0491 | -27.8646
18.0503 | -35.0496 | -27.8725
18.0564 | -35.0502 | -27.88
18.0625 | -35.0507 | -27.8877
18.0686 | -35.0512 | -27.8957
18.0747 | -35.0517 | -27.9035
18.0808 | -35.0523 | -27.9113
18.0869 | -35.0528 | -27.9191
18.093 | -35.0533 | -27.9269
18.0991 | -35.0538 | -27.9348
18.1052 | -35.0544 | -27.9423
18.1113 | -35.0549 | -27.9501
18.1174 | -35.0554 | -27.9577
18.1236 | -35.0559 | -27.9655
18.1297 | -35.0565 | -27.9734
18.1358 | -35.057 | -27.981
18.1419 | -35.0575 | -27.9889
18.148 | -35.058 | -27.9965
18.1541 | -35.0586 | -28.0042
18.1602 | -35.0591 | -28.0119
18.1663 | -35.0596 | -28.0195
18.1724 | -35.0601 | -28.0273
18.1785 | -35.0607 | -28.0352
18.1846 | -35.0612 | -28.0428
18.1907 | -35.0616 | -28.0507
18.1968 | -35.0621 | -28.0585
18.2029 | -35.0626 | -28.0662
18.209 | -35.0632 | -28.074
18.2151 | -35.0637 | -28.0819
18.2212 | -35.0642 | -28.0896
18.2273 | -35.0647 | -28.0975
18.2334 | -35.0653 | -28.1055
18.2395 | -35.0657 | -28.1131
18.2456 | -35.0663 | -28.1209
18.2517 | -35.0668 | -28.1288
18.2578 | -35.0672 | -28.1366
18.2639 | -35.0678 | -28.1444
18.2701 | -35.0683 | -28.1521
18.2762 | -35.0688 | -28.1597
18.2823 | -35.0693 | -28.1678
18.2884 | -35.0698 | -28.1755
18.2945 | -35.0703 | -28.1831
18.3006 | -35.0709 | -28.1904
18.3067 | -35.0714 | -28.1991
18.3128 | -35.072 | -28.2064
18.3189 | -35.0725 | -28.2139
18.325 | -35.073 | -28.2215
18.3311 | -35.0735 | -28.2293
18.3372 | -35.0739 | -28.2371
18.3433 | -35.0745 | -28.2443
18.3494 | -35.0751 | -28.2525
18.3555 | -35.0757 | -28.26
18.3616 | -35.0762 | -28.2675
18.3677 | -35.0767 | -28.2758
18.3738 | -35.0771 | -28.2837
18.3799 | -35.0777 | -28.2915
18.386 | -35.0783 | -28.2991
18.3921 | -35.0788 | -28.3067
18.3982 | -35.0793 | -28.3143
18.4043 | -35.0798 | -28.322
18.4105 | -35.0803 | -28.3302
18.4166 | -35.0808 | -28.3382
18.4227 | -35.0813 | -28.346
18.4288 | -35.0818 | -28.3537
18.4349 | -35.0823 | -28.3612
18.441 | -35.0828 | -28.3688
18.4471 | -35.0833 | -28.3764
18.4532 | -35.0839 | -28.3842
18.4593 | -35.0844 | -28.3916
18.4654 | -35.0849 | -28.3993
18.4715 | -35.0854 | -28.4073
18.4776 | -35.0859 | -28.415
18.4837 | -35.0865 | -28.423
18.4898 | -35.087 | -28.4306
18.4959 | -35.0876 | -28.4386
18.502 | -35.0881 | -28.4463
18.5081 | -35.0886 | -28.4539
18.5142 | -35.0891 | -28.4616
18.5203 | -35.0896 | -28.4689
18.5264 | -35.0901 | -28.4765
18.5325 | -35.0906 | -28.4844
18.5386 | -35.0912 | -28.4918
18.5447 | -35.0917 | -28.4994
18.5508 | -35.0922 | -28.5071
18.557 | -35.0928 | -28.5154
18.5631 | -35.0933 | -28.523
18.5692 | -35.0938 | -28.5304
18.5753 | -35.0943 | -28.538
18.5814 | -35.0948 | -28.5453
18.5875 | -35.0953 | -28.5531
18.5936 | -35.0958 | -28.5609
18.5997 | -35.0963 | -28.5685
18.6058 | -35.0968 | -28.5761
18.6119 | -35.0973 | -28.5838
18.618 | -35.0979 | -28.5915
18.6241 | -35.0984 | -28.5995
18.6302 | -35.0988 | -28.6072
18.6363 | -35.0993 | -28.6145
18.6424 | -35.0998 | -28.622
18.6485 | -35.1003 | -28.6298
18.6546 | -35.1008 | -28.6373
18.6607 | -35.1013 | -28.6449
18.6668 | -35.1018 | -28.653
18.6729 | -35.1024 | -28.6608
18.679 | -35.1029 | -28.6683
18.6851 | -35.1034 | -28.6757
18.6912 | -35.1039 | -28.6832
18.6974 | -35.1044 | -28.6905
18.7035 | -35.1049 | -28.6981
18.7096 | -35.1054 | -28.7059
18.7157 | -35.1059 | -28.7133
18.7218 | -35.1064 | -28.7208
18.7279 | -35.1068 | -28.7285
18.734 | -35.1073 | -28.7364
18.7401 | -35.1078 | -28.7438
18.7462 | -35.1083 | -28.751
18.7523 | -35.1088 | -28.7584
18.7584 | -35.1093 | -28.7661
18.7645 | -35.1097 | -28.7736
18.7706 | -35.1102 | -28.7814
18.7767 | -35.1107 | -28.7888
18.7828 | -35.1111 | -28.7963
18.7889 | -35.1116 | -28.804
18.795 | -35.1122 | -28.8119
18.8011 | -35.1126 | -28.8197
18.8072 | -35.1132 | -28.8277
18.8133 | -35.1137 | -28.8356
18.8194 | -35.1141 | -28.843
18.8255 | -35.1146 | -28.8506
18.8316 | -35.1152 | -28.8579
18.8377 | -35.1157 | -28.8656
18.8439 | -35.1162 | -28.8735
18.85 | -35.1167 | -28.8807
18.8561 | -35.1172 | -28.8878
18.8622 | -35.1176 | -28.8955
18.8683 | -35.1182 | -28.9031
18.8744 | -35.1186 | -28.9105
18.8805 | -35.1191 | -28.918
18.8866 | -35.1196 | -28.926
18.8927 | -35.1201 | -28.9334
18.8988 | -35.1206 | -28.9412
18.9049 | -35.121 | -28.9492
18.911 | -35.1215 | -28.9569
18.9171 | -35.122 | -28.964
18.9232 | -35.1225 | -28.9715
18.9293 | -35.123 | -28.9792
18.9354 | -35.1234 | -28.9866
18.9415 | -35.1239 | -28.9939
18.9476 | -35.1243 | -29.0025
18.9537 | -35.1246 | -29.0098
18.9598 | -35.1251 | -29.0169
18.9659 | -35.1256 | -29.0245
18.972 | -35.1262 | -29.0315
18.9781 | -35.1265 | -29.0396
18.9843 | -35.1269 | -29.0457
18.9904 | -35.1273 | -29.0536
18.9965 | -35.1277 | -29.0608
19.0026 | -35.1282 | -29.0685
19.0087 | -35.1288 | -29.0756
19.0148 | -35.1292 | -29.0829
19.0209 | -35.1297 | -29.091
19.027 | -35.1303 | -29.099
19.0331 | -35.1308 | -29.1064
19.0392 | -35.1313 | -29.1137
19.0453 | -35.1317 | -29.1206
19.0514 | -35.1323 | -29.1286
19.0575 | -35.1328 | -29.1358
19.0636 | -35.1332 | -29.1431
19.0697 | -35.1338 | -29.1507
19.0758 | -35.1343 | -29.1586
19.0819 | -35.1348 | -29.1658
19.088 | -35.1352 | -29.1736
19.0941 | -35.1357 | -29.1805
19.1002 | -35.1361 | -29.188
19.1063 | -35.1366 | -29.1955
19.1124 | -35.1371 | -29.2028
19.1185 | -35.1375 | -29.2103
19.1246 | -35.138 | -29.2181
19.1308 | -35.1386 | -29.226
19.1369 | -35.139 | -29.2331
19.143 | -35.1396 | -29.2403
19.1491 | -35.1401 | -29.2473
19.1552 | -35.1406 | -29.2555
19.1613 | -35.1411 | -29.2631
19.1674 | -35.1416 | -29.2705
19.1735 | -35.1421 | -29.2779
19.1796 | -35.1424 | -29.2851
19.1857 | -35.143 | -29.2924
19.1918 | -35.1434 | -29.3003
19.1979 | -35.1439 | -29.3071
19.204 | -35.1444 | -29.3146
19.2101 | -35.1449 | -29.3222
19.2162 | -35.1454 | -29.3297
19.2223 | -35.146 | -29.337
19.2284 | -35.1464 | -29.3453
19.2345 | -35.1469 | -29.3532
19.2406 | -35.1474 | -29.36
19.2467 | -35.1478 | -29.3678
19.2528 | -35.1483 | -29.3753
19.2589 | -35.1488 | -29.3825
19.265 | -35.1493 | -29.3906
19.2712 | -35.1498 | -29.3985
19.2773 | -35.1503 | -29.4058
19.2834 | -35.1507 | -29.4134
19.2895 | -35.1512 | -29.4209
19.2956 | -35.1517 | -29.4285
19.3017 | -35.1522 | -29.4363
19.3078 | -35.1527 | -29.4434
19.3139 | -35.1532 | -29.4505
19.32 | -35.1536 | -29.4576
19.3261 | -35.1541 | -29.4648
19.3322 | -35.1546 | -29.4719
19.3383 | -35.1551 | -29.4791
19.3444 | -35.1555 | -29.4861
19.3505 | -35.1561 | -29.4935
19.3566 | -35.1566 | -29.5012
19.3627 | -35.1571 | -29.5086
19.3688 | -35.1576 | -29.5162
19.3749 | -35.1581 | -29.5234
19.381 | -35.1586 | -29.5302
19.3871 | -35.1591 | -29.5374
19.3932 | -35.1595 | -29.5449
19.3993 | -35.1599 | -29.5522
19.4054 | -35.1604 | -29.5601
19.4115 | -35.1608 | -29.5684
19.4177 | -35.1613 | -29.5755
19.4238 | -35.1617 | -29.5832
19.4299 | -35.1622 | -29.59
19.436 | -35.1627 | -29.5978
19.4421 | -35.1632 | -29.6054
19.4482 | -35.1636 | -29.6125
19.4543 | -35.1641 | -29.6197
19.4604 | -35.1646 | -29.627
19.4665 | -35.1651 | -29.6338
19.4726 | -35.1656 | -29.641
19.4787 | -35.1661 | -29.6484
19.4848 | -35.1665 | -29.6553
19.4909 | -35.1669 | -29.6625
19.497 | -35.1674 | -29.6698
19.5031 | -35.1678 | -29.6769
19.5092 | -35.1682 | -29.6847
19.5153 | -35.1687 | -29.6919
19.5214 | -35.1691 | -29.6992
19.5275 | -35.1696 | -29.7064
19.5336 | -35.17 | -29.7143
19.5397 | -35.1705 | -29.7216
19.5458 | -35.1708 | -29.7289
19.5519 | -35.1713 | -29.7364
19.5581 | -35.1717 | -29.7434
19.5642 | -35.1722 | -29.7506
19.5703 | -35.1726 | -29.7582
19.5764 | -35.1731 | -29.7657
19.5825 | -35.1737 | -29.7736
19.5886 | -35.1741 | -29.7812
19.5947 | -35.1746 | -29.7883
19.6008 | -35.175 | -29.795
19.6069 | -35.1756 | -29.8027
19.613 | -35.1761 | -29.8094
19.6191 | -35.1765 | -29.8167
19.6252 | -35.177 | -29.8242
19.6313 | -35.1775 | -29.8314
19.6374 | -35.178 | -29.8379
19.6435 | -35.1785 | -29.8451
19.6496 | -35.179 | -29.8528
19.6557 | -35.1795 | -29.8605
19.6618 | -35.18 | -29.8672
19.6679 | -35.1805 | -29.8751
19.674 | -35.1808 | -29.8829
19.6801 | -35.1812 | -29.8896
19.6862 | -35.1815 | -29.8966
19.6923 | -35.182 | -29.9038
19.6984 | -35.1824 | -29.9107
19.7046 | -35.1829 | -29.9184
19.7107 | -35.1835 | -29.9263
19.7168 | -35.1839 | -29.9339
19.7229 | -35.1845 | -29.9418
19.729 | -35.1848 | -29.9488
19.7351 | -35.1853 | -29.9559
19.7412 | -35.1858 | -29.9635
19.7473 | -35.1861 | -29.9709
19.7534 | -35.1866 | -29.9782
19.7595 | -35.187 | -29.9857
19.7656 | -35.1875 | -29.9929
19.7717 | -35.1879 | -30
19.7778 | -35.1883 | -30.0067
19.7839 | -35.1888 | -30.0139
19.79 | -35.1894 | -30.0212
19.7961 | -35.1897 | -30.0287
19.8022 | -35.1902 | -30.0356
19.8083 | -35.1907 | -30.043
19.8144 | -35.1911 | -30.0501
19.8205 | -35.1916 | -30.0572
19.8266 | -35.1921 | -30.0647
19.8327 | -35.1926 | -30.0713
19.8388 | -35.193 | -30.0782
19.845 | -35.1934 | -30.0857
19.8511 | -35.1939 | -30.0932
19.8572 | -35.1942 | -30.0997
19.8633 | -35.1947 | -30.1067
19.8694 | -35.1952 | -30.1144
19.8755 | -35.1956 | -30.1214
19.8816 | -35.1961 | -30.1291
19.8877 | -35.1965 | -30.1364
19.8938 | -35.1969 | -30.1437
19.8999 | -35.1974 | -30.1508
19.906 | -35.1978 | -30.1583
19.9121 | -35.1983 | -30.1651
19.9182 | -35.1987 | -30.1719
19.9243 | -35.1991 | -30.1795
19.9304 | -35.1994 | -30.1864
19.9365 | -35.1998 | -30.1934
19.9426 | -35.2003 | -30.1999
19.9487 | -35.2007 | -30.2072
19.9548 | -35.2013 | -30.2141
19.9609 | -35.2017 | -30.2223
19.967 | -35.2021 | -30.2297
19.9731 | -35.2026 | -30.2366
19.9792 | -35.203 | -30.2434
19.9853 | -35.2035 | -30.2505
19.9915 | -35.204 | -30.2575
19.9976 | -35.2045 | -30.2649
20.0037 | -35.2049 | -30.2729
20.0098 | -35.2054 | -30.2803
20.0159 | -35.2057 | -30.2869
20.022 | -35.2062 | -30.2941
20.0281 | -35.2066 | -30.3017
20.0342 | -35.207 | -30.3087
20.0403 | -35.2075 | -30.3164
20.0464 | -35.2081 | -30.3232
20.0525 | -35.2085 | -30.3303
20.0586 | -35.2089 | -30.3375
20.0647 | -35.2093 | -30.3451
20.0708 | -35.2097 | -30.3522
20.0769 | -35.2102 | -30.3602
20.083 | -35.2107 | -30.3667
20.0891 | -35.2111 | -30.3736
20.0952 | -35.2115 | -30.38
20.1013 | -35.212 | -30.3873
20.1074 | -35.2125 | -30.3944
20.1135 | -35.2129 | -30.401
20.1196 | -35.2133 | -30.4073
20.1257 | -35.2137 | -30.4144
20.1319 | -35.2141 | -30.4227
20.138 | -35.2146 | -30.43
20.1441 | -35.2151 | -30.4399
20.1502 | -35.2154 | -30.4489
20.1563 | -35.2159 | -30.4557
20.1624 | -35.2164 | -30.4623
20.1685 | -35.2169 | -30.4686
20.1746 | -35.2174 | -30.4748
20.1807 | -35.2178 | -30.4834
20.1868 | -35.2183 | -30.4919
20.1929 | -35.2188 | -30.4992
20.199 | -35.2192 | -30.5064
20.2051 | -35.2196 | -30.5133
20.2112 | -35.2202 | -30.52
20.2173 | -35.2207 | -30.5274
20.2234 | -35.2211 | -30.5345
20.2295 | -35.2215 | -30.5411
20.2356 | -35.2218 | -30.548
20.2417 | -35.2222 | -30.5551
20.2478 | -35.2226 | -30.5618
20.2539 | -35.223 | -30.5689
20.26 | -35.2235 | -30.5755
20.2661 | -35.2239 | -30.5825
20.2723 | -35.2243 | -30.5896
20.2784 | -35.2246 | -30.596
20.2845 | -35.225 | -30.6029
20.2906 | -35.2254 | -30.6097
20.2967 | -35.2259 | -30.6165
20.3028 | -35.2263 | -30.6226
20.3089 | -35.2267 | -30.6296
20.315 | -35.2271 | -30.6365
20.3211 | -35.2275 | -30.6438
20.3272 | -35.228 | -30.6507
20.3333 | -35.2285 | -30.6576
20.3394 | -35.2289 | -30.6645
20.3455 | -35.2293 | -30.6713
20.3516 | -35.2298 | -30.678
20.3577 | -35.2301 | -30.685
20.3638 | -35.2305 | -30.6917
20.3699 | -35.2309 | -30.6981
20.376 | -35.2313 | -30.705
20.3821 | -35.2318 | -30.712
20.3882 | -35.2322 | -30.7189
20.3943 | -35.2327 | -30.7254
20.4004 | -35.2331 | -30.7317
20.4065 | -35.2336 | -30.7387
20.4126 | -35.2341 | -30.7458
20.4188 | -35.2346 | -30.7524
20.4249 | -35.2351 | -30.7593
20.431 | -35.2355 | -30.766
20.4371 | -35.2359 | -30.7728
20.4432 | -35.2363 | -30.7798
20.4493 | -35.2366 | -30.7866
20.4554 | -35.2371 | -30.7935
20.4615 | -35.2376 | -30.8002
20.4676 | -35.238 | -30.8074
20.4737 | -35.2385 | -30.8143
20.4798 | -35.2389 | -30.8213
20.4859 | -35.2393 | -30.8282
20.492 | -35.2397 | -30.8354
20.4981 | -35.2402 | -30.8423
20.5042 | -35.2407 | -30.8493
20.5103 | -35.241 | -30.8559
20.5164 | -35.2414 | -30.8624
20.5225 | -35.2418 | -30.869
20.5286 | -35.2422 | -30.8759
20.5347 | -35.2427 | -30.8825
20.5408 | -35.2431 | -30.8891
20.5469 | -35.2435 | -30.8959
20.553 | -35.244 | -30.9029
20.5592 | -35.2444 | -30.9098
20.5653 | -35.2449 | -30.9165
20.5714 | -35.2454 | -30.9229
20.5775 | -35.2458 | -30.9294
20.5836 | -35.2462 | -30.9361
20.5897 | -35.2466 | -30.9434
20.5958 | -35.2471 | -30.9501
20.6019 | -35.2475 | -30.957
20.608 | -35.2479 | -30.9644
20.6141 | -35.2483 | -30.9708
20.6202 | -35.2488 | -30.9777
20.6263 | -35.2492 | -30.9841
20.6324 | -35.2496 | -30.9904
20.6385 | -35.2501 | -30.9972
20.6446 | -35.2506 | -31.0039
20.6507 | -35.251 | -31.0112
20.6568 | -35.2514 | -31.0178
20.6629 | -35.2518 | -31.0242
20.669 | -35.2523 | -31.0314
20.6751 | -35.2526 | -31.0378
20.6812 | -35.2531 | -31.0441
20.6873 | -35.2535 | -31.0506
20.6934 | -35.2538 | -31.0574
20.6995 | -35.2542 | -31.0639
20.7057 | -35.2546 | -31.0708
20.7118 | -35.2551 | -31.0779
20.7179 | -35.2554 | -31.0841
20.724 | -35.2558 | -31.0908
20.7301 | -35.2563 | -31.0972
20.7362 | -35.2566 | -31.104
20.7423 | -35.257 | -31.1101
20.7484 | -35.2575 | -31.1171
20.7545 | -35.2578 | -31.1235
20.7606 | -35.2583 | -31.1304
20.7667 | -35.2587 | -31.1371
20.7728 | -35.2591 | -31.1439
20.7789 | -35.2595 | -31.1506
20.785 | -35.2599 | -31.157
20.7911 | -35.2604 | -31.1639
20.7972 | -35.2608 | -31.1707
20.8033 | -35.2612 | -31.1775
20.8094 | -35.2616 | -31.1841
20.8155 | -35.262 | -31.1909
20.8216 | -35.2625 | -31.1973
20.8277 | -35.263 | -31.2034
20.8338 | -35.2634 | -31.2098
20.8399 | -35.2639 | -31.2165
20.8461 | -35.2642 | -31.2237
20.8522 | -35.2645 | -31.2299
20.8583 | -35.265 | -31.236
20.8644 | -35.2654 | -31.242
20.8705 | -35.2659 | -31.2496
20.8766 | -35.2663 | -31.2553
20.8827 | -35.2666 | -31.2626
20.8888 | -35.267 | -31.2695
20.8949 | -35.2673 | -31.2743
20.901 | -35.2677 | -31.2811
20.9071 | -35.2682 | -31.2873
20.9132 | -35.2686 | -31.2937
20.9193 | -35.269 | -31.3
20.9254 | -35.2695 | -31.3064
20.9315 | -35.27 | -31.3123
20.9376 | -35.2704 | -31.3187
20.9437 | -35.2709 | -31.3256
20.9498 | -35.2714 | -31.332
20.9559 | -35.2718 | -31.3389
20.962 | -35.2722 | -31.3454
20.9681 | -35.2727 | -31.3515
20.9742 | -35.2731 | -31.358
20.9803 | -35.2735 | -31.365
20.9864 | -35.274 | -31.3715
20.9926 | -35.2744 | -31.3784
20.9987 | -35.2748 | -31.385
21.0048 | -35.2753 | -31.3913
21.0109 | -35.2757 | -31.3975
21.017 | -35.2762 | -31.4041
21.0231 | -35.2766 | -31.4109
21.0292 | -35.2771 | -31.4175
21.0353 | -35.2775 | -31.423
21.0414 | -35.2779 | -31.429
21.0475 | -35.2784 | -31.4356
21.0536 | -35.2789 | -31.4425
21.0597 | -35.2793 | -31.4492
21.0658 | -35.2796 | -31.4555
21.0719 | -35.28 | -31.4613
21.078 | -35.2804 | -31.4675
21.0841 | -35.2808 | -31.474
21.0902 | -35.2812 | -31.4807
21.0963 | -35.2817 | -31.4876
21.1024 | -35.2821 | -31.4942
21.1085 | -35.2824 | -31.5008
21.1146 | -35.2828 | -31.508
21.1207 | -35.2832 | -31.5151
21.1268 | -35.2836 | -31.522
21.133 | -35.2841 | -31.5284
21.1391 | -35.2844 | -31.5345
21.1452 | -35.2849 | -31.5411
21.1513 | -35.2853 | -31.5479
21.1574 | -35.2857 | -31.5545
21.1635 | -35.2861 | -31.5611
21.1696 | -35.2865 | -31.5675
21.1757 | -35.2869 | -31.5726
21.1818 | -35.2873 | -31.5794
21.1879 | -35.2878 | -31.5856
21.194 | -35.2883 | -31.5923
21.2001 | -35.2888 | -31.5991
21.2062 | -35.2892 | -31.605
21.2123 | -35.2895 | -31.6118
21.2184 | -35.29 | -31.6177
21.2245 | -35.2903 | -31.6249
21.2306 | -35.2907 | -31.6315
21.2367 | -35.2911 | -31.638
21.2428 | -35.2915 | -31.6447
21.2489 | -35.2917 | -31.6507
21.255 | -35.2921 | -31.6572
21.2611 | -35.2925 | -31.6639
21.2672 | -35.2929 | -31.6702
21.2733 | -35.2933 | -31.6776
21.2795 | -35.2936 | -31.6857
21.2856 | -35.2942 | -31.6925
21.2917 | -35.2946 | -31.6992
21.2978 | -35.2949 | -31.7055
21.3039 | -35.2955 | -31.7117
21.31 | -35.296 | -31.7185
21.3161 | -35.2964 | -31.7253
21.3222 | -35.2968 | -31.7325
21.3283 | -35.2973 | -31.7386
21.3344 | -35.2977 | -31.7449
21.3405 | -35.2981 | -31.7518
21.3466 | -35.2986 | -31.7567
21.3527 | -35.299 | -31.7634
21.3588 | -35.2996 | -31.7695
21.3649 | -35.3002 | -31.7753
21.371 | -35.3007 | -31.7819
21.3771 | -35.3012 | -31.7887
21.3832 | -35.3015 | -31.7947
21.3893 | -35.302 | -31.8013
21.3954 | -35.3024 | -31.8085
21.4015 | -35.303 | -31.8152
21.4076 | -35.3034 | -31.8211
21.4137 | -35.3039 | -31.8278
21.4199 | -35.3043 | -31.834
21.426 | -35.3048 | -31.8392
21.4321 | -35.3053 | -31.846
21.4382 | -35.3058 | -31.8523
21.4443 | -35.3062 | -31.8587
21.4504 | -35.3066 | -31.8646
21.4565 | -35.3071 | -31.8713
21.4626 | -35.3075 | -31.8784
21.4687 | -35.3078 | -31.8851
21.4748 | -35.3082 | -31.8917
21.4809 | -35.3084 | -31.8981
21.487 | -35.3089 | -31.9047
21.4931 | -35.3091 | -31.9098
21.4992 | -35.3094 | -31.9159
21.5053 | -35.3099 | -31.922
21.5114 | -35.3103 | -31.9281
21.5175 | -35.3107 | -31.9342
21.5236 | -35.3112 | -31.9404
21.5297 | -35.3116 | -31.9464
21.5358 | -35.312 | -31.9533
21.5419 | -35.3125 | -31.9593
21.548 | -35.3129 | -31.9658
21.5541 | -35.3133 | -31.973
21.5602 | -35.3137 | -31.9801
21.5664 | -35.3142 | -31.9869
21.5725 | -35.3146 | -31.9931
21.5786 | -35.3151 | -31.9993
21.5847 | -35.3155 | -32.0058
21.5908 | -35.316 | -32.0129
21.5969 | -35.3163 | -32.0198
21.603 | -35.3167 | -32.0265
21.6091 | -35.317 | -32.0331
21.6152 | -35.3175 | -32.0397
21.6213 | -35.318 | -32.0468
21.6274 | -35.3185 | -32.0524
21.6335 | -35.319 | -32.058
21.6396 | -35.3193 | -32.0645
21.6457 | -35.3197 | -32.0724
21.6518 | -35.3201 | -32.0782
21.6579 | -35.3204 | -32.0847
21.664 | -35.3208 | -32.091
21.6701 | -35.3214 | -32.0961
21.6762 | -35.3216 | -32.1021
21.6823 | -35.322 | -32.1083
21.6884 | -35.3224 | -32.1143
21.6945 | -35.3229 | -32.1227
21.7006 | -35.3236 | -32.1312
21.7068 | -35.3242 | -32.135
21.7129 | -35.3248 | -32.1417
21.719 | -35.3251 | -32.1496
21.7251 | -35.3257 | -32.1559
21.7312 | -35.3261 | -32.1625
21.7373 | -35.3267 | -32.1672
21.7434 | -35.3271 | -32.1729
21.7495 | -35.3277 | -32.1789
21.7556 | -35.3281 | -32.1847
21.7617 | -35.3285 | -32.1913
21.7678 | -35.3289 | -32.198
21.7739 | -35.3292 | -32.2048
21.78 | -35.3296 | -32.2109
21.7861 | -35.3301 | -32.2182
21.7922 | -35.3305 | -32.2238
21.7983 | -35.3309 | -32.2301
21.8044 | -35.3313 | -32.2362
21.8105 | -35.3317 | -32.2429
21.8166 | -35.332 | -32.2501
21.8227 | -35.3324 | -32.2567
21.8288 | -35.3328 | -32.2635
21.8349 | -35.3332 | -32.2691
21.841 | -35.3337 | -32.2757
21.8471 | -35.3341 | -32.2809
21.8533 | -35.3345 | -32.2879
21.8594 | -35.3348 | -32.294
21.8655 | -35.3351 | -32.3005
21.8716 | -35.3353 | -32.3065
21.8777 | -35.3359 | -32.3127
21.8838 | -35.3363 | -32.3191
21.8899 | -35.3367 | -32.3257
21.896 | -35.3371 | -32.3327
21.9021 | -35.3375 | -32.3391
21.9082 | -35.3379 | -32.3466
21.9143 | -35.3385 | -32.3534
21.9204 | -35.3389 | -32.3589
21.9265 | -35.3393 | -32.3655
21.9326 | -35.3398 | -32.3721
21.9387 | -35.3402 | -32.3782
21.9448 | -35.3405 | -32.3848
21.9509 | -35.3409 | -32.3925
21.957 | -35.3414 | -32.399
21.9631 | -35.3418 | -32.4049
21.9692 | -35.3421 | -32.4106
21.9753 | -35.3426 | -32.4177
21.9814 | -35.3429 | -32.4246
21.9875 | -35.3434 | -32.4298
21.9937 | -35.3437 | -32.4363
21.9998 | -35.3442 | -32.4436
22.0059 | -35.3447 | -32.4501
22.012 | -35.3451 | -32.4561
22.0181 | -35.3457 | -32.4618
22.0242 | -35.346 | -32.4684
22.0303 | -35.3464 | -32.4743
22.0364 | -35.3468 | -32.4809
22.0425 | -35.3473 | -32.488
22.0486 | -35.3477 | -32.4945
22.0547 | -35.3482 | -32.5003
22.0608 | -35.3487 | -32.507
22.0669 | -35.3491 | -32.5136
22.073 | -35.3494 | -32.5201
22.0791 | -35.3497 | -32.5264
22.0852 | -35.3502 | -32.5332
22.0913 | -35.3506 | -32.5402
22.0974 | -35.351 | -32.5471
22.1035 | -35.3514 | -32.5538
22.1096 | -35.3518 | -32.5596
22.1157 | -35.3523 | -32.5664
22.1218 | -35.3527 | -32.5728
22.1279 | -35.3531 | -32.5791
22.134 | -35.3534 | -32.5857
22.1402 | -35.3538 | -32.5913
22.1463 | -35.3543 | -32.5985
22.1524 | -35.3546 | -32.6047
22.1585 | -35.355 | -32.6105
22.1646 | -35.3555 | -32.6158
22.1707 | -35.356 | -32.6218
22.1768 | -35.3566 | -32.628
22.1829 | -35.3571 | -32.6346
22.189 | -35.3575 | -32.6415
22.1951 | -35.3579 | -32.6472
22.2012 | -35.3582 | -32.6534
22.2073 | -35.3587 | -32.6598
22.2134 | -35.3592 | -32.6658
22.2195 | -35.3596 | -32.6715
22.2256 | -35.3601 | -32.6784
22.2317 | -35.3605 | -32.6853
22.2378 | -35.3611 | -32.6919
22.2439 | -35.3614 | -32.6983
22.25 | -35.3618 | -32.7036
22.2561 | -35.3622 | -32.7095
22.2622 | -35.3625 | -32.7165
22.2683 | -35.363 | -32.7221
22.2744 | -35.3635 | -32.7284
22.2806 | -35.3638 | -32.7345
22.2867 | -35.3643 | -32.7401
22.2928 | -35.3649 | -32.7464
22.2989 | -35.3655 | -32.7536
22.305 | -35.3659 | -32.7595
22.3111 | -35.3663 | -32.7662
22.3172 | -35.3668 | -32.7732
22.3233 | -35.3671 | -32.7795
22.3294 | -35.3675 | -32.7855
22.3355 | -35.3679 | -32.7914
22.3416 | -35.3682 | -32.7987
22.3477 | -35.3686 | -32.8049
22.3538 | -35.369 | -32.8114
22.3599 | -35.3694 | -32.8181
22.366 | -35.3698 | -32.8249
22.3721 | -35.3705 | -32.8294
22.3782 | -35.3708 | -32.8364
22.3843 | -35.3713 | -32.843
22.3904 | -35.3718 | -32.8494
22.3965 | -35.3723 | -32.856
22.4026 | -35.3727 | -32.8615
22.4087 | -35.3732 | -32.8662
22.4148 | -35.3735 | -32.8724
22.4209 | -35.3739 | -32.8786
22.4271 | -35.3743 | -32.885
22.4332 | -35.3749 | -32.8908
22.4393 | -35.3754 | -32.8973
22.4454 | -35.3757 | -32.9041
22.4515 | -35.3761 | -32.9098
22.4576 | -35.3766 | -32.9164
22.4637 | -35.3769 | -32.9229
22.4698 | -35.3772 | -32.9284
22.4759 | -35.3775 | -32.9342
22.482 | -35.378 | -32.9406
22.4881 | -35.3784 | -32.9463
22.4942 | -35.3791 | -32.9518
22.5003 | -35.3795 | -32.9584
22.5064 | -35.38 | -32.9646
22.5125 | -35.3804 | -32.9707
22.5186 | -35.3808 | -32.975
22.5247 | -35.3807 | -32.9833
22.5308 | -35.3808 | -32.9902
22.5369 | -35.3812 | -32.9962
22.543 | -35.3815 | -33.0002
22.5491 | -35.3822 | -33.0097
22.5552 | -35.3828 | -33.0171
22.5613 | -35.3832 | -33.0227
22.5675 | -35.3835 | -33.0296
22.5736 | -35.3839 | -33.0363
22.5797 | -35.3844 | -33.0428
22.5858 | -35.3847 | -33.0496
22.5919 | -35.3851 | -33.0571
22.598 | -35.3854 | -33.0635
22.6041 | -35.3859 | -33.0689
22.6102 | -35.3863 | -33.0749
22.6163 | -35.3866 | -33.0799
22.6224 | -35.387 | -33.0865
22.6285 | -35.3874 | -33.093
22.6346 | -35.3878 | -33.0991
22.6407 | -35.3882 | -33.1057
22.6468 | -35.3887 | -33.113
22.6529 | -35.389 | -33.1203
22.659 | -35.3894 | -33.1268
22.6651 | -35.3898 | -33.1328
22.6712 | -35.3902 | -33.1404
22.6773 | -35.3908 | -33.1461
22.6834 | -35.3912 | -33.1528
22.6895 | -35.3916 | -33.1595
22.6956 | -35.3919 | -33.1662
22.7017 | -35.3923 | -33.1716
22.7079 | -35.3927 | -33.1782
22.714 | -35.3931 | -33.1832
22.7201 | -35.3935 | -33.1892
22.7262 | -35.394 | -33.1954
22.7323 | -35.3946 | -33.2018
22.7384 | -35.395 | -33.208
22.7445 | -35.3956 | -33.215
22.7506 | -35.396 | -33.2217
22.7567 | -35.3964 | -33.228
22.7628 | -35.3967 | -33.2349
22.7689 | -35.3972 | -33.2415
22.775 | -35.3977 | -33.2476
22.7811 | -35.3981 | -33.2542
22.7872 | -35.3985 | -33.2604
22.7933 | -35.3987 | -33.2663
22.7994 | -35.3991 | -33.2725
22.8055 | -35.3995 | -33.2788
22.8116 | -35.3999 | -33.2856
22.8177 | -35.4006 | -33.2911
22.8238 | -35.4009 | -33.2971
22.8299 | -35.4015 | -33.303
22.836 | -35.4019 | -33.3098
22.8421 | -35.4024 | -33.3166
22.8482 | -35.4029 | -33.3239
22.8544 | -35.4033 | -33.3304
22.8605 | -35.4037 | -33.3366
22.8666 | -35.404 | -33.3421
22.8727 | -35.4046 | -33.3484
22.8788 | -35.4049 | -33.3537
22.8849 | -35.4053 | -33.3602
22.891 | -35.4056 | -33.3663
22.8971 | -35.4059 | -33.3726
22.9032 | -35.4064 | -33.3776
22.9093 | -35.4068 | -33.3837
22.9154 | -35.4072 | -33.3906
22.9215 | -35.4076 | -33.3974
22.9276 | -35.4081 | -33.4031
22.9337 | -35.4086 | -33.4094
22.9398 | -35.4091 | -33.4136
22.9459 | -35.4095 | -33.42
22.952 | -35.4099 | -33.426
22.9581 | -35.4104 | -33.4322
22.9642 | -35.4108 | -33.4389
22.9703 | -35.4112 | -33.4457
22.9764 | -35.4116 | -33.4519
22.9825 | -35.4122 | -33.4583
22.9886 | -35.4126 | -33.4651
22.9948 | -35.413 | -33.4702
23.0009 | -35.4134 | -33.4771
23.007 | -35.4139 | -33.4824
23.0131 | -35.4144 | -33.4888
23.0192 | -35.4148 | -33.4945
23.0253 | -35.4152 | -33.5006
23.0314 | -35.4156 | -33.507
23.0375 | -35.416 | -33.513
23.0436 | -35.4163 | -33.5198
23.0497 | -35.4168 | -33.5263
23.0558 | -35.4172 | -33.5325
23.0619 | -35.4176 | -33.539
23.068 | -35.4182 | -33.5449
23.0741 | -35.4185 | -33.5507
23.0802 | -35.4189 | -33.5569
23.0863 | -35.4193 | -33.5631
23.0924 | -35.4198 | -33.5695
23.0985 | -35.4205 | -33.5753
23.1046 | -35.4209 | -33.5818
23.1107 | -35.4213 | -33.5893
23.1168 | -35.4218 | -33.5955
23.1229 | -35.4221 | -33.6007
23.129 | -35.4226 | -33.6061
23.1351 | -35.4231 | -33.6122
23.1413 | -35.4234 | -33.6185
23.1474 | -35.4241 | -33.6242
23.1535 | -35.4246 | -33.6305
23.1596 | -35.425 | -33.637
23.1657 | -35.4255 | -33.6441
23.1718 | -35.426 | -33.6501
23.1779 | -35.4264 | -33.6569
23.184 | -35.4268 | -33.6636
23.1901 | -35.4273 | -33.6703
23.1962 | -35.4277 | -33.6768
23.2023 | -35.4279 | -33.6825
23.2084 | -35.4284 | -33.6882
23.2145 | -35.429 | -33.6949
23.2206 | -35.4294 | -33.7011
23.2267 | -35.4301 | -33.7074
23.2328 | -35.4305 | -33.7135
23.2389 | -35.4308 | -33.7207
23.245 | -35.4311 | -33.7281
23.2511 | -35.4316 | -33.7344
23.2572 | -35.4321 | -33.7407
23.2633 | -35.4326 | -33.7463
23.2694 | -35.433 | -33.7524
23.2755 | -35.4336 | -33.7593
23.2817 | -35.4339 | -33.766
23.2878 | -35.4344 | -33.7721
23.2939 | -35.4347 | -33.7778
23.3 | -35.4352 | -33.7851
23.3061 | -35.4356 | -33.7909
23.3122 | -35.4361 | -33.7969
23.3183 | -35.4365 | -33.803
23.3244 | -35.4371 | -33.8095
23.3305 | -35.4376 | -33.8156
23.3366 | -35.438 | -33.8222
23.3427 | -35.4384 | -33.8291
23.3488 | -35.4389 | -33.8349
23.3549 | -35.4392 | -33.8425
23.361 | -35.4397 | -33.849
23.3671 | -35.4404 | -33.8578
23.3732 | -35.4407 | -33.8667
23.3793 | -35.441 | -33.8747
23.3854 | -35.4414 | -33.8807
23.3915 | -35.4417 | -33.883
23.3976 | -35.4422 | -33.8895
23.4037 | -35.4426 | -33.8959
23.4098 | -35.4431 | -33.9022
23.4159 | -35.4435 | -33.9078
23.422 | -35.4441 | -33.9132
23.4282 | -35.4448 | -33.9193
23.4343 | -35.4453 | -33.9261
23.4404 | -35.4457 | -33.9336
23.4465 | -35.446 | -33.9389
23.4526 | -35.4462 | -33.9446
23.4587 | -35.4468 | -33.9507
23.4648 | -35.4472 | -33.9577
23.4709 | -35.4477 | -33.9631
23.477 | -35.4481 | -33.9698
23.4831 | -35.4485 | -33.9761
23.4892 | -35.449 | -33.9823
23.4953 | -35.4493 | -33.9888
23.5014 | -35.4498 | -33.9938
23.5075 | -35.4502 | -34.0002
23.5136 | -35.4507 | -34.0055
23.5197 | -35.4512 | -34.0122
23.5258 | -35.4517 | -34.0192
23.5319 | -35.4522 | -34.0258
23.538 | -35.4525 | -34.0319
23.5441 | -35.4532 | -34.0389
23.5502 | -35.4535 | -34.0443
23.5563 | -35.454 | -34.0504
23.5624 | -35.4544 | -34.0571
23.5686 | -35.4549 | -34.0634
23.5747 | -35.4554 | -34.0708
23.5808 | -35.4557 | -34.0768
23.5869 | -35.4563 | -34.0826
23.593 | -35.4569 | -34.0903
23.5991 | -35.4572 | -34.0967
23.6052 | -35.4575 | -34.1029
23.6113 | -35.4579 | -34.1091
23.6174 | -35.4584 | -34.1146
23.6235 | -35.4589 | -34.1219
23.6296 | -35.4593 | -34.1283
23.6357 | -35.4598 | -34.1348
23.6418 | -35.4603 | -34.1407
23.6479 | -35.4607 | -34.1457
23.654 | -35.4612 | -34.1521
23.6601 | -35.4617 | -34.158
23.6662 | -35.4621 | -34.1636
23.6723 | -35.4626 | -34.1693
23.6784 | -35.463 | -34.1759
23.6845 | -35.4635 | -34.1824
23.6906 | -35.464 | -34.189
23.6967 | -35.4644 | -34.1955
23.7028 | -35.4648 | -34.2014
23.7089 | -35.4651 | -34.2084
23.7151 | -35.4656 | -34.2147
23.7212 | -35.4661 | -34.2214
23.7273 | -35.4664 | -34.2271
23.7334 | -35.4669 | -34.2343
23.7395 | -35.4674 | -34.2405
23.7456 | -35.4677 | -34.2452
23.7517 | -35.4682 | -34.2513
23.7578 | -35.4686 | -34.2575
23.7639 | -35.4691 | -34.2641
23.77 | -35.4695 | -34.2713
23.7761 | -35.4699 | -34.2776
23.7822 | -35.4703 | -34.2835
23.7883 | -35.4708 | -34.2898
23.7944 | -35.4713 | -34.2955
23.8005 | -35.4718 | -34.3016
23.8066 | -35.4721 | -34.3077
23.8127 | -35.4724 | -34.3138
23.8188 | -35.4729 | -34.3201
23.8249 | -35.4734 | -34.3269
23.831 | -35.4738 | -34.3337
23.8371 | -35.4743 | -34.3398
23.8432 | -35.4746 | -34.3461
23.8493 | -35.4748 | -34.3524
23.8555 | -35.4753 | -34.3586
23.8616 | -35.4758 | -34.3652
23.8677 | -35.4762 | -34.3717
23.8738 | -35.4766 | -34.3782
23.8799 | -35.4771 | -34.3849
23.886 | -35.4775 | -34.3918
23.8921 | -35.478 | -34.3985
23.8982 | -35.4783 | -34.4046
23.9043 | -35.4788 | -34.4115
23.9104 | -35.4792 | -34.4172
23.9165 | -35.4796 | -34.4239
23.9226 | -35.48 | -34.4294
23.9287 | -35.4805 | -34.4357
23.9348 | -35.481 | -34.4421
23.9409 | -35.4815 | -34.4484
23.947 | -35.4819 | -34.4552
23.9531 | -35.4824 | -34.4609
23.9592 | -35.4828 | -34.4677
23.9653 | -35.4833 | -34.4741
23.9714 | -35.4836 | -34.4799
23.9775 | -35.4841 | -34.4861
23.9836 | -35.4845 | -34.4928
23.9897 | -35.4849 | -34.4991
23.9958 | -35.4853 | -34.5055
24.002 | -35.4856 | -34.5117
24.0081 | -35.4861 | -34.5184
24.0142 | -35.4865 | -34.5245
24.0203 | -35.487 | -34.5309
24.0264 | -35.4876 | -34.5371
24.0325 | -35.488 | -34.5428
24.0386 | -35.4883 | -34.5491
24.0447 | -35.4887 | -34.5558
24.0508 | -35.4891 | -34.5617
24.0569 | -35.4896 | -34.568
24.063 | -35.49 | -34.5745
24.0691 | -35.4905 | -34.5806
24.0752 | -35.4908 | -34.5862
24.0813 | -35.4913 | -34.5926
24.0874 | -35.4917 | -34.5988
24.0935 | -35.4922 | -34.6056
24.0996 | -35.4926 | -34.6119
24.1057 | -35.493 | -34.6183
24.1118 | -35.4934 | -34.6247
24.1179 | -35.4938 | -34.6315
24.124 | -35.4942 | -34.6376
24.1301 | -35.4945 | -34.6437
24.1362 | -35.495 | -34.6502
24.1424 | -35.4955 | -34.6569
24.1485 | -35.496 | -34.6629
24.1546 | -35.4964 | -34.6694
24.1607 | -35.4969 | -34.6756
24.1668 | -35.4974 | -34.6814
24.1729 | -35.4978 | -34.6877
24.179 | -35.4983 | -34.6935
24.1851 | -35.4987 | -34.7
24.1912 | -35.499 | -34.7053
24.1973 | -35.4994 | -34.7112
24.2034 | -35.4999 | -34.7156
24.2095 | -35.5003 | -34.7216
24.2156 | -35.5009 | -34.728
24.2217 | -35.5012 | -34.7333
24.2278 | -35.5016 | -34.7383
24.2339 | -35.502 | -34.7442
24.24 | -35.5024 | -34.7499
24.2461 | -35.5029 | -34.7559
24.2522 | -35.5033 | -34.7619
24.2583 | -35.5037 | -34.7679
24.2644 | -35.5041 | -34.7746
24.2705 | -35.5046 | -34.781
24.2766 | -35.5049 | -34.7871
24.2827 | -35.5054 | -34.793
24.2889 | -35.5058 | -34.7994
24.295 | -35.5063 | -34.8055
24.3011 | -35.5068 | -34.8111
24.3072 | -35.5073 | -34.8171
24.3133 | -35.5077 | -34.8229
24.3194 | -35.5081 | -34.8287
24.3255 | -35.5086 | -34.8352
24.3316 | -35.5091 | -34.8416
24.3377 | -35.5095 | -34.8473
24.3438 | -35.5099 | -34.8534
24.3499 | -35.5102 | -34.8598
24.356 | -35.5106 | -34.866
24.3621 | -35.511 | -34.8722
24.3682 | -35.5115 | -34.8784
24.3743 | -35.5118 | -34.8843
24.3804 | -35.5123 | -34.89
24.3865 | -35.5127 | -34.896
24.3926 | -35.513 | -34.9024
24.3987 | -35.5135 | -34.9081
24.4048 | -35.5138 | -34.9142
24.4109 | -35.5143 | -34.9199
24.417 | -35.5147 | -34.9268
24.4231 | -35.5152 | -34.9337
24.4293 | -35.5158 | -34.9399
24.4354 | -35.5162 | -34.9462
24.4415 | -35.5167 | -34.9525
24.4476 | -35.517 | -34.9587
24.4537 | -35.5174 | -34.9652
24.4598 | -35.5178 | -34.9719
24.4659 | -35.5181 | -34.9778
24.472 | -35.5186 | -34.9839
24.4781 | -35.519 | -34.99
24.4842 | -35.5193 | -34.9962
24.4903 | -35.5199 | -35.0031
24.4964 | -35.5203 | -35.0099
24.5025 | -35.5207 | -35.0162
24.5086 | -35.5212 | -35.0226
24.5147 | -35.5215 | -35.0286
24.5208 | -35.5219 | -35.0342
24.5269 | -35.5223 | -35.0402
24.533 | -35.5227 | -35.0462
24.5391 | -35.5231 | -35.0521
24.5452 | -35.5237 | -35.0575
24.5513 | -35.5242 | -35.0641
24.5574 | -35.5246 | -35.0704
24.5635 | -35.525 | -35.0765
24.5696 | -35.5254 | -35.0832
24.5758 | -35.5259 | -35.0894
24.5819 | -35.5263 | -35.0954
24.588 | -35.5267 | -35.1016
24.5941 | -35.5271 | -35.107
24.6002 | -35.5275 | -35.1131
24.6063 | -35.5279 | -35.1195
24.6124 | -35.5283 | -35.1252
24.6185 | -35.5287 | -35.1316
24.6246 | -35.5291 | -35.1367
24.6307 | -35.5295 | -35.1431
24.6368 | -35.5299 | -35.1496
24.6429 | -35.5304 | -35.1564
24.649 | -35.5309 | -35.1627
24.6551 | -35.5312 | -35.169
24.6612 | -35.5317 | -35.1738
24.6673 | -35.532 | -35.1799
24.6734 | -35.5325 | -35.1862
24.6795 | -35.5329 | -35.1925
24.6856 | -35.5333 | -35.1988
24.6917 | -35.5337 | -35.2052
24.6978 | -35.5341 | -35.2109
24.7039 | -35.5345 | -35.2172
24.71 | -35.5349 | -35.2232
24.7162 | -35.5354 | -35.2298
24.7223 | -35.5357 | -35.2359
24.7284 | -35.5361 | -35.2417
24.7345 | -35.5365 | -35.2478
24.7406 | -35.537 | -35.2543
24.7467 | -35.5375 | -35.2609
24.7528 | -35.5379 | -35.2673
24.7589 | -35.5383 | -35.2735
24.765 | -35.5387 | -35.2796
24.7711 | -35.5392 | -35.2858
24.7772 | -35.5396 | -35.2922
24.7833 | -35.54 | -35.2985
24.7894 | -35.5404 | -35.3048
24.7955 | -35.5408 | -35.3109
24.8016 | -35.5412 | -35.3177
24.8077 | -35.5415 | -35.3237
24.8138 | -35.542 | -35.3301
24.8199 | -35.5424 | -35.336
24.826 | -35.5428 | -35.3423
24.8321 | -35.5433 | -35.3484
24.8382 | -35.5436 | -35.3539
24.8443 | -35.544 | -35.36
24.8504 | -35.5444 | -35.3661
24.8565 | -35.5448 | -35.3724
24.8627 | -35.5452 | -35.3785
24.8688 | -35.5457 | -35.3846
24.8749 | -35.5461 | -35.3909
24.881 | -35.5465 | -35.3971
24.8871 | -35.547 | -35.4035
24.8932 | -35.5475 | -35.4094
24.8993 | -35.548 | -35.4162
24.9054 | -35.5484 | -35.4226
24.9115 | -35.5488 | -35.4285
24.9176 | -35.5492 | -35.4353
24.9237 | -35.5496 | -35.4414
24.9298 | -35.55 | -35.4474
24.9359 | -35.5504 | -35.4533
24.942 | -35.5507 | -35.4588
24.9481 | -35.5511 | -35.4646
24.9542 | -35.5515 | -35.4708
24.9603 | -35.552 | -35.477
24.9664 | -35.5524 | -35.483
24.9725 | -35.5527 | -35.4892
24.9786 | -35.5532 | -35.4952
24.9847 | -35.5536 | -35.5013
24.9908 | -35.554 | -35.5076
24.9969 | -35.5543 | -35.5133
25.0031 | -35.5548 | -35.5195
25.0092 | -35.5552 | -35.5254
25.0153 | -35.5556 | -35.5314
25.0214 | -35.5559 | -35.5376
25.0275 | -35.5563 | -35.5439
25.0336 | -35.5567 | -35.55
25.0397 | -35.5571 | -35.5559
25.0458 | -35.5577 | -35.5614
25.0519 | -35.5581 | -35.5671
25.058 | -35.5585 | -35.5731
25.0641 | -35.5588 | -35.5791
25.0702 | -35.5593 | -35.5851
25.0763 | -35.5597 | -35.5914
25.0824 | -35.5601 | -35.5969
25.0885 | -35.5605 | -35.603
25.0946 | -35.5609 | -35.6091
25.1007 | -35.5613 | -35.6148
25.1068 | -35.5617 | -35.6205
25.1129 | -35.5622 | -35.6267
25.119 | -35.5626 | -35.6322
25.1251 | -35.5631 | -35.6381
25.1312 | -35.5635 | -35.6443
25.1373 | -35.5639 | -35.6504
25.1435 | -35.5643 | -35.6564
25.1496 | -35.5648 | -35.6615
25.1557 | -35.5652 | -35.6675
25.1618 | -35.5657 | -35.6735
25.1679 | -35.5661 | -35.6798
25.174 | -35.5665 | -35.6855
25.1801 | -35.5669 | -35.6916
25.1862 | -35.5673 | -35.697
25.1923 | -35.5677 | -35.7034
25.1984 | -35.5681 | -35.7093
25.2045 | -35.5685 | -35.715
25.2106 | -35.5689 | -35.7209
25.2167 | -35.5693 | -35.7272
25.2228 | -35.5696 | -35.7328
25.2289 | -35.57 | -35.739
25.235 | -35.5705 | -35.7451
25.2411 | -35.5709 | -35.751
25.2472 | -35.5713 | -35.7569
25.2533 | -35.5717 | -35.7627
25.2594 | -35.5721 | -35.7684
25.2655 | -35.5726 | -35.7748
25.2716 | -35.573 | -35.7807
25.2777 | -35.5734 | -35.7867
25.2838 | -35.5738 | -35.7926
25.29 | -35.5742 | -35.7983
25.2961 | -35.5746 | -35.8039
25.3022 | -35.575 | -35.8095
25.3083 | -35.5754 | -35.8155
25.3144 | -35.5758 | -35.8214
25.3205 | -35.5763 | -35.8274
25.3266 | -35.5767 | -35.8329
25.3327 | -35.5771 | -35.8389
25.3388 | -35.5775 | -35.8449
25.3449 | -35.5779 | -35.851
25.351 | -35.5784 | -35.8563
25.3571 | -35.5788 | -35.8624
25.3632 | -35.5792 | -35.8681
25.3693 | -35.5796 | -35.8743
25.3754 | -35.58 | -35.8801
25.3815 | -35.5803 | -35.886
25.3876 | -35.5808 | -35.8921
25.3937 | -35.5812 | -35.898
25.3998 | -35.5815 | -35.9035
25.4059 | -35.582 | -35.9094
25.412 | -35.5823 | -35.9153
25.4181 | -35.5827 | -35.9212
25.4242 | -35.5831 | -35.9275
25.4304 | -35.5835 | -35.9332
25.4365 | -35.584 | -35.9387
25.4426 | -35.5843 | -35.9444
25.4487 | -35.5848 | -35.9505
25.4548 | -35.5852 | -35.9561
25.4609 | -35.5856 | -35.962
25.467 | -35.586 | -35.9678
25.4731 | -35.5864 | -35.9736
25.4792 | -35.5868 | -35.9794
25.4853 | -35.5872 | -35.9852
25.4914 | -35.5875 | -35.991
25.4975 | -35.5879 | -35.9963
25.5036 | -35.5883 | -36.0017
25.5097 | -35.5887 | -36.0072
25.5158 | -35.5892 | -36.013
25.5219 | -35.5896 | -36.0191
25.528 | -35.59 | -36.025
25.5341 | -35.5904 | -36.0308
25.5402 | -35.5907 | -36.0368
25.5463 | -35.5912 | -36.0425
25.5524 | -35.5916 | -36.0484
25.5585 | -35.592 | -36.0542
25.5646 | -35.5924 | -36.0601
25.5707 | -35.5928 | -36.0658
25.5769 | -35.5932 | -36.0712
25.583 | -35.5936 | -36.0768
25.5891 | -35.594 | -36.0827
25.5952 | -35.5944 | -36.0884
25.6013 | -35.5949 | -36.094
25.6074 | -35.5953 | -36.0993
25.6135 | -35.5957 | -36.1054
25.6196 | -35.5962 | -36.1113
25.6257 | -35.5966 | -36.1173
25.6318 | -35.597 | -36.1229
25.6379 | -35.5974 | -36.1285
25.644 | -35.5977 | -36.1346
25.6501 | -35.5981 | -36.1403
25.6562 | -35.5984 | -36.1463
25.6623 | -35.5988 | -36.1519
25.6684 | -35.5992 | -36.1574
25.6745 | -35.5995 | -36.1633
25.6806 | -35.5999 | -36.1686
25.6867 | -35.6003 | -36.1746
25.6928 | -35.6007 | -36.1802
25.6989 | -35.601 | -36.1859
25.705 | -35.6014 | -36.1917
25.7111 | -35.6018 | -36.1973
25.7173 | -35.6022 | -36.2024
25.7234 | -35.6025 | -36.2084
25.7295 | -35.6029 | -36.2139
25.7356 | -35.6033 | -36.2198
25.7417 | -35.6037 | -36.2253
25.7478 | -35.6041 | -36.2309
25.7539 | -35.6045 | -36.2365
25.76 | -35.6049 | -36.2424
25.7661 | -35.6053 | -36.248
25.7722 | -35.6057 | -36.2539
25.7783 | -35.6061 | -36.2597
25.7844 | -35.6065 | -36.2651
25.7905 | -35.6068 | -36.2707
25.7966 | -35.6072 | -36.2758
25.8027 | -35.6076 | -36.2814
25.8088 | -35.608 | -36.2875
25.8149 | -35.6084 | -36.2931
25.821 | -35.6087 | -36.2983
25.8271 | -35.6091 | -36.3041
25.8332 | -35.6095 | -36.3103
25.8393 | -35.6099 | -36.3159
25.8454 | -35.6103 | -36.3218
25.8515 | -35.6106 | -36.3274
25.8576 | -35.611 | -36.3328
25.8638 | -35.6113 | -36.3383
25.8699 | -35.6118 | -36.3442
25.876 | -35.6122 | -36.3498
25.8821 | -35.6125 | -36.3553
25.8882 | -35.6129 | -36.3604
25.8943 | -35.6133 | -36.366
25.9004 | -35.6137 | -36.3714
25.9065 | -35.6141 | -36.3771
25.9126 | -35.6145 | -36.3829
25.9187 | -35.6149 | -36.3883
25.9248 | -35.6153 | -36.3937
25.9309 | -35.6157 | -36.3995
25.937 | -35.616 | -36.4051
25.9431 | -35.6165 | -36.4108
25.9492 | -35.6169 | -36.4166
25.9553 | -35.6173 | -36.4223
25.9614 | -35.6177 | -36.4279
25.9675 | -35.6181 | -36.4336
25.9736 | -35.6185 | -36.4393
25.9797 | -35.6189 | -36.4452
25.9858 | -35.6193 | -36.4507
25.9919 | -35.6196 | -36.4568
25.998 | -35.6199 | -36.4626
26.0042 | -35.6203 | -36.4684
26.0103 | -35.6207 | -36.4739
26.0164 | -35.6211 | -36.4793
26.0225 | -35.6215 | -36.4852
26.0286 | -35.6219 | -36.4908
26.0347 | -35.6222 | -36.496
26.0408 | -35.6226 | -36.5018
26.0469 | -35.6229 | -36.5072
26.053 | -35.6232 | -36.5131
26.0591 | -35.6236 | -36.5187
26.0652 | -35.6239 | -36.5238
26.0713 | -35.6243 | -36.5292
26.0774 | -35.6246 | -36.5345
26.0835 | -35.625 | -36.5403
26.0896 | -35.6254 | -36.5459
26.0957 | -35.6258 | -36.5517
26.1018 | -35.6262 | -36.5569
26.1079 | -35.6266 | -36.5625
26.114 | -35.6269 | -36.5683
26.1201 | -35.6273 | -36.5738
26.1262 | -35.6277 | -36.5792
26.1323 | -35.6281 | -36.5846
26.1384 | -35.6285 | -36.59
26.1445 | -35.6288 | -36.5958
26.1507 | -35.6292 | -36.6015
26.1568 | -35.6296 | -36.607
26.1629 | -35.63 | -36.6128
26.169 | -35.6304 | -36.6184
26.1751 | -35.6308 | -36.6236
26.1812 | -35.6311 | -36.6291
26.1873 | -35.6315 | -36.6346
26.1934 | -35.6319 | -36.64
26.1995 | -35.6323 | -36.6458
26.2056 | -35.6326 | -36.6509
26.2117 | -35.633 | -36.6563
26.2178 | -35.6334 | -36.6618
26.2239 | -35.6337 | -36.6678
26.23 | -35.6341 | -36.6735
26.2361 | -35.6344 | -36.6791
26.2422 | -35.6348 | -36.6839
26.2483 | -35.6352 | -36.6893
26.2544 | -35.6355 | -36.6947
26.2605 | -35.6358 | -36.7002
26.2666 | -35.6362 | -36.706
26.2727 | -35.6366 | -36.7112
26.2788 | -35.637 | -36.7166
26.2849 | -35.6374 | -36.7219
26.2911 | -35.6378 | -36.7269
26.2972 | -35.6382 | -36.7325
26.3033 | -35.6385 | -36.738
26.3094 | -35.6389 | -36.7438
26.3155 | -35.6393 | -36.7488
26.3216 | -35.6396 | -36.7547
26.3277 | -35.64 | -36.7601
26.3338 | -35.6403 | -36.7657
26.3399 | -35.6407 | -36.7713
26.346 | -35.641 | -36.7766
26.3521 | -35.6414 | -36.7821
26.3582 | -35.6418 | -36.7875
26.3643 | -35.6422 | -36.793
26.3704 | -35.6425 | -36.7986
26.3765 | -35.6429 | -36.8039
26.3826 | -35.6432 | -36.809
26.3887 | -35.6436 | -36.8145
26.3948 | -35.644 | -36.8201
26.4009 | -35.6443 | -36.8258
26.407 | -35.6448 | -36.8311
26.4131 | -35.6451 | -36.8364
26.4192 | -35.6455 | -36.8412
26.4253 | -35.6459 | -36.8469
26.4314 | -35.6463 | -36.8525
26.4376 | -35.6467 | -36.8576
26.4437 | -35.647 | -36.863
26.4498 | -35.6474 | -36.8685
26.4559 | -35.6477 | -36.8737
26.462 | -35.6481 | -36.8793
26.4681 | -35.6484 | -36.8853
26.4742 | -35.6488 | -36.8908
26.4803 | -35.6492 | -36.8963
26.4864 | -35.6495 | -36.9015
26.4925 | -35.6499 | -36.9067
26.4986 | -35.6503 | -36.9118
26.5047 | -35.6507 | -36.9174
26.5108 | -35.651 | -36.9231
26.5169 | -35.6513 | -36.9283
26.523 | -35.6516 | -36.9333
26.5291 | -35.6519 | -36.9387
26.5352 | -35.6523 | -36.9439
26.5413 | -35.6527 | -36.9495
26.5474 | -35.6531 | -36.9546
26.5535 | -35.6534 | -36.9602
26.5596 | -35.6538 | -36.9651
26.5657 | -35.6541 | -36.9708
26.5718 | -35.6545 | -36.9761
26.578 | -35.6548 | -36.9815
26.5841 | -35.6551 | -36.9868
26.5902 | -35.6555 | -36.9921
26.5963 | -35.6559 | -36.9968
26.6024 | -35.6563 | -37.0023
26.6085 | -35.6566 | -37.0076
26.6146 | -35.657 | -37.0131
26.6207 | -35.6574 | -37.0182
26.6268 | -35.6578 | -37.0236
26.6329 | -35.6581 | -37.029
26.639 | -35.6584 | -37.0345
26.6451 | -35.6588 | -37.0401
26.6512 | -35.6592 | -37.0456
26.6573 | -35.6595 | -37.0511
26.6634 | -35.6598 | -37.0563
26.6695 | -35.6602 | -37.0618
26.6756 | -35.6605 | -37.0672
26.6817 | -35.6609 | -37.0729
26.6878 | -35.6613 | -37.0783
26.6939 | -35.6617 | -37.0838
26.7 | -35.662 | -37.0887
26.7061 | -35.6623 | -37.0941
26.7122 | -35.6627 | -37.0992
26.7183 | -35.663 | -37.1045
26.7245 | -35.6633 | -37.1093
26.7306 | -35.6636 | -37.1147
26.7367 | -35.664 | -37.1194
26.7428 | -35.6644 | -37.1246
26.7489 | -35.6648 | -37.13
26.755 | -35.6651 | -37.1356
26.7611 | -35.6654 | -37.1407
26.7672 | -35.6658 | -37.1459
26.7733 | -35.6661 | -37.1515
26.7794 | -35.6665 | -37.157
26.7855 | -35.6669 | -37.1629
26.7916 | -35.6672 | -37.1678
26.7977 | -35.6676 | -37.1733
26.8038 | -35.6679 | -37.1786
26.8099 | -35.6682 | -37.1842
26.816 | -35.6685 | -37.1894
26.8221 | -35.6689 | -37.1946
26.8282 | -35.6692 | -37.2
26.8343 | -35.6696 | -37.2051
26.8404 | -35.67 | -37.2098
26.8465 | -35.6703 | -37.2152
26.8526 | -35.6707 | -37.2206
26.8587 | -35.6711 | -37.2257
26.8649 | -35.6714 | -37.231
26.871 | -35.6718 | -37.2363
26.8771 | -35.672 | -37.2412
26.8832 | -35.6724 | -37.2466
26.8893 | -35.6727 | -37.2517
26.8954 | -35.6731 | -37.257
26.9015 | -35.6734 | -37.2624
26.9076 | -35.6738 | -37.2678
26.9137 | -35.6742 | -37.2727
26.9198 | -35.6745 | -37.2775
26.9259 | -35.6748 | -37.2828
26.932 | -35.6752 | -37.2881
26.9381 | -35.6755 | -37.2934
26.9442 | -35.6758 | -37.2985
26.9503 | -35.6761 | -37.3035
26.9564 | -35.6765 | -37.3088
26.9625 | -35.6768 | -37.314
26.9686 | -35.6772 | -37.3192
26.9747 | -35.6775 | -37.3242
26.9808 | -35.6779 | -37.3289
26.9869 | -35.6782 | -37.3343
26.993 | -35.6786 | -37.3397
26.9991 | -35.6789 | -37.3449
27.0052 | -35.6793 | -37.3504
27.0114 | -35.6796 | -37.3555
27.0175 | -35.6799 | -37.3605
27.0236 | -35.6803 | -37.3656
27.0297 | -35.6806 | -37.3709
27.0358 | -35.6809 | -37.3762
27.0419 | -35.6812 | -37.3812
27.048 | -35.6815 | -37.3859
27.0541 | -35.6818 | -37.3906
27.0602 | -35.6822 | -37.3961
27.0663 | -35.6825 | -37.4016
27.0724 | -35.6828 | -37.4069
27.0785 | -35.6832 | -37.4124
27.0846 | -35.6835 | -37.4174
27.0907 | -35.6838 | -37.4226
27.0968 | -35.6841 | -37.4277
27.1029 | -35.6845 | -37.4327
27.109 | -35.6848 | -37.438
27.1151 | -35.6852 | -37.4432
27.1212 | -35.6855 | -37.4481
27.1273 | -35.6858 | -37.4532
27.1334 | -35.6861 | -37.4584
27.1395 | -35.6864 | -37.4633
27.1456 | -35.6867 | -37.4688
27.1518 | -35.687 | -37.4741
27.1579 | -35.6873 | -37.4791
27.164 | -35.6876 | -37.4844
27.1701 | -35.6879 | -37.4894
27.1762 | -35.6882 | -37.4948
27.1823 | -35.6885 | -37.4998
27.1884 | -35.6888 | -37.5051
27.1945 | -35.6891 | -37.5102
27.2006 | -35.6894 | -37.5156
27.2067 | -35.6897 | -37.5208
27.2128 | -35.69 | -37.5257
27.2189 | -35.6903 | -37.531
27.225 | -35.6906 | -37.536
27.2311 | -35.691 | -37.5409
27.2372 | -35.6913 | -37.546
27.2433 | -35.6916 | -37.5507
27.2494 | -35.692 | -37.556
27.2555 | -35.6923 | -37.5612
27.2616 | -35.6926 | -37.566
27.2677 | -35.6929 | -37.5711
27.2738 | -35.6933 | -37.5765
27.2799 | -35.6936 | -37.5813
27.286 | -35.6939 | -37.5866
27.2921 | -35.6942 | -37.5921
27.2983 | -35.6945 | -37.5971
27.3044 | -35.6948 | -37.6022
27.3105 | -35.695 | -37.6073
27.3166 | -35.6954 | -37.6122
27.3227 | -35.6957 | -37.6172
27.3288 | -35.696 | -37.6222
27.3349 | -35.6963 | -37.6271
27.341 | -35.6966 | -37.6323
27.3471 | -35.697 | -37.6373
27.3532 | -35.6973 | -37.6425
27.3593 | -35.6976 | -37.6477
27.3654 | -35.6979 | -37.653
27.3715 | -35.6982 | -37.6581
27.3776 | -35.6986 | -37.6634
27.3837 | -35.6989 | -37.6681
27.3898 | -35.6992 | -37.6731
27.3959 | -35.6996 | -37.6784
27.402 | -35.6999 | -37.6829
27.4081 | -35.7002 | -37.6878
27.4142 | -35.7005 | -37.6927
27.4203 | -35.7008 | -37.6979
27.4264 | -35.7011 | -37.703
27.4325 | -35.7014 | -37.7083
27.4387 | -35.7017 | -37.7131
27.4448 | -35.702 | -37.7181
27.4509 | -35.7022 | -37.7234
27.457 | -35.7026 | -37.7286
27.4631 | -35.7029 | -37.7338
27.4692 | -35.7032 | -37.7389
27.4753 | -35.7035 | -37.7435
27.4814 | -35.7038 | -37.7486
27.4875 | -35.7041 | -37.7538
27.4936 | -35.7044 | -37.7586
27.4997 | -35.7047 | -37.7638
27.5058 | -35.705 | -37.7687
27.5119 | -35.7053 | -37.7734
27.518 | -35.7056 | -37.7785
27.5241 | -35.7059 | -37.7835
27.5302 | -35.7061 | -37.7887
27.5363 | -35.7065 | -37.7938
27.5424 | -35.7068 | -37.7986
27.5485 | -35.7071 | -37.8034
27.5546 | -35.7074 | -37.8084
27.5607 | -35.7077 | -37.8132
27.5668 | -35.708 | -37.8182
27.5729 | -35.7083 | -37.8229
27.5791 | -35.7086 | -37.8276
27.5852 | -35.7088 | -37.8325
27.5913 | -35.7092 | -37.8376
27.5974 | -35.7094 | -37.843
27.6035 | -35.7097 | -37.848
27.6096 | -35.7101 | -37.8533
27.6157 | -35.7104 | -37.858
27.6218 | -35.7107 | -37.8631
27.6279 | -35.711 | -37.8681
27.634 | -35.7113 | -37.8729
27.6401 | -35.7116 | -37.8779
27.6462 | -35.7118 | -37.8828
27.6523 | -35.7121 | -37.8873
27.6584 | -35.7124 | -37.8926
27.6645 | -35.7127 | -37.8978
27.6706 | -35.713 | -37.9028
27.6767 | -35.7134 | -37.9079
27.6828 | -35.7136 | -37.9127
27.6889 | -35.7139 | -37.9177
27.695 | -35.7142 | -37.9226
27.7011 | -35.7145 | -37.9276
27.7072 | -35.7148 | -37.9326
27.7133 | -35.7151 | -37.9374
27.7194 | -35.7154 | -37.9422
27.7256 | -35.7157 | -37.9472
27.7317 | -35.7159 | -37.9521
27.7378 | -35.7162 | -37.9572
27.7439 | -35.7165 | -37.9623
27.75 | -35.7168 | -37.9672
27.7561 | -35.717 | -37.9719
27.7622 | -35.7173 | -37.977
27.7683 | -35.7176 | -37.9818
27.7744 | -35.718 | -37.9867
27.7805 | -35.7183 | -37.9916
27.7866 | -35.7186 | -37.9965
27.7927 | -35.7188 | -38.0013
27.7988 | -35.7191 | -38.0065
27.8049 | -35.7194 | -38.0114
27.811 | -35.7197 | -38.0165
27.8171 | -35.7201 | -38.0214
27.8232 | -35.7203 | -38.0258
27.8293 | -35.7206 | -38.0308
27.8354 | -35.7209 | -38.0356
27.8415 | -35.7211 | -38.0405
27.8476 | -35.7214 | -38.0457
27.8537 | -35.7217 | -38.0506
27.8598 | -35.722 | -38.0552
27.866 | -35.7222 | -38.0603
27.8721 | -35.7226 | -38.0652
27.8782 | -35.7228 | -38.0704
27.8843 | -35.7231 | -38.0752
27.8904 | -35.7234 | -38.0802
27.8965 | -35.7236 | -38.0848
27.9026 | -35.7239 | -38.0896
27.9087 | -35.7242 | -38.0947
27.9148 | -35.7245 | -38.0995
27.9209 | -35.7248 | -38.1045
27.927 | -35.725 | -38.1096
27.9331 | -35.7253 | -38.1146
27.9392 | -35.7255 | -38.1197
27.9453 | -35.7258 | -38.1242
27.9514 | -35.7261 | -38.1295
27.9575 | -35.7263 | -38.1346
27.9636 | -35.7266 | -38.1395
27.9697 | -35.7269 | -38.1442
27.9758 | -35.7272 | -38.1491
27.9819 | -35.7274 | -38.1539
27.988 | -35.7277 | -38.1588
27.9941 | -35.728 | -38.1636
28.0002 | -35.7282 | -38.1689
28.0063 | -35.7285 | -38.1738
28.0125 | -35.7288 | -38.1786
28.0186 | -35.7291 | -38.1837
28.0247 | -35.7294 | -38.1882
28.0308 | -35.7297 | -38.193
28.0369 | -35.7299 | -38.1977
28.043 | -35.7302 | -38.2025
28.0491 | -35.7304 | -38.2073
28.0552 | -35.7307 | -38.2122
28.0613 | -35.731 | -38.2172
28.0674 | -35.7313 | -38.2224
28.0735 | -35.7316 | -38.2266

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4503 | -43.0167
21.4908 | -35.5272 | -34.9976
16.4948 | -34.9183 | -27.6161
12.6602 | -34.5615 | -21.3516
9.7171 | -34.3478 | -16.4515
7.4582 | -34.1985 | -12.4855
5.7244 | -34.0953 | -9.2644
4.3936 | -34.0295 | -6.679
3.3722 | -34.0153 | -4.6381
2.5883 | -34.0141 | -3.0926
1.9866 | -33.9967 | -1.8561
1.5248 | -33.9993 | -0.86541
1.1703 | -34.0019 | -0.22177
1 | -34.0113 | 0.099006
0.89822 | -33.9932 | 0.36947
0.68941 | -33.9887 | 0.94317
0.52919 | -33.9986 | 1.3872
0.40612 | -33.9833 | 1.9957
0.31171 | -33.9699 | 2.4467
0.23928 | -33.985 | 3.162
0.18358 | -33.9804 | 3.5784
0.14091 | -34.0096 | 4.4488
0.10819 | -34.014 | 5.6073
0.083017 | -34.0473 | 7.0688
0.063705 | -34.0834 | 8.9284
0.048923 | -34.1339 | 11.4124
0.037479 | -34.2446 | 14.587
0.028801 | -34.4254 | 18.5868
0.022078 | -34.7269 | 23.8833
0.017023 | -35.1891 | 30.2471
0.013018 | -35.9869 | 37.9076
0.010014 | -37.088 | 45.7892

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003673 | -10.1847
0.007347 | -26.435
0.01102 | -34.9442
0.014693 | -40.8179
0.018366 | -43.2686
0.02204 | -44.8822
0.025713 | -47.4634
0.029386 | -48.3613
0.033059 | -49.4993
0.036733 | -50.6845
0.040406 | -51.4921
0.044079 | -53.5408
0.047752 | -54.8916
0.051426 | -56.008
0.055099 | -56.283
0.058772 | -56.4072
0.062445 | -56.8218
0.066119 | -57.8697
0.069792 | -60.0085
0.073465 | -59.198
0.077139 | -59.5183
0.080812 | -59.8555
0.084485 | -60.5776
0.088158 | -62.5091
0.091832 | -62.6406
0.095505 | -63.5975
0.099178 | -63.7761
0.10285 | -63.2317
0.10652 | -63.8764
0.1102 | -65.3559
0.11387 | -65.6816
0.11754 | -65.8753
0.12122 | -66.851
0.12489 | -66.6899
0.12856 | -66.9901
0.13224 | -67.239
0.13591 | -67.4067
0.13958 | -67.0776
0.14326 | -67.0882
0.14693 | -68.0808
0.1506 | -68.6996
0.15428 | -68.5214
0.15795 | -68.5177
0.16162 | -68.0096
0.1653 | -68.6986
0.16897 | -69.1554
0.17264 | -69.0401
0.17632 | -69.1551
0.17999 | -69.8404
0.18366 | -70.4967
0.18734 | -70.2353
0.19101 | -70.6473
0.19468 | -70.3903
0.19836 | -70.6411
0.20203 | -70.9341
0.2057 | -71.474
0.20938 | -72.3322
0.21305 | -71.7313
0.21672 | -71.8456
0.2204 | -72.0611
0.22407 | -72.3949
0.22774 | -72.1001
0.23142 | -72.0476
0.23509 | -72.9089
0.23876 | -73.5393
0.24244 | -73.6236
0.24611 | -73.76
0.24978 | -74.6089
0.25345 | -74.876
0.25713 | -74.4244
0.2608 | -74.0013
0.26448 | -73.5519
0.26815 | -73.7879
0.27182 | -74.0403
0.27549 | -74.1212
0.27917 | -74.6798
0.28284 | -74.6869
0.28652 | -74.447
0.29019 | -74.3168
0.29386 | -75.1238
0.29753 | -75.3656
0.30121 | -74.5892
0.30488 | -75.1604
0.30855 | -76.6791
0.31223 | -75.4129
0.3159 | -74.5879
0.31957 | -76.2121
0.32325 | -76.9268
0.32692 | -77.1202
0.33059 | -77.1181
0.33427 | -77.0978
0.33794 | -76.3091
0.34161 | -75.5867
0.34529 | -75.8119
0.34896 | -75.6202
0.35263 | -76.256
0.35631 | -76.87
0.35998 | -77.7682
0.36365 | -77.7463
0.36733 | -77.531
0.371 | -76.7617
0.37467 | -76.5071
0.37835 | -76.5976
0.38202 | -76.3983
0.38569 | -77.4727
0.38937 | -77.8612
0.39304 | -78.6642
0.39671 | -78.8612
0.40039 | -78.8299
0.40406 | -78.7696
0.40773 | -78.3543
0.41141 | -78.9612
0.41508 | -79.263
0.41875 | -79.1145
0.42242 | -79.451
0.4261 | -79.3844
0.42977 | -79.8834
0.43345 | -79.4565
0.43712 | -79.2438
0.44079 | -80.0381
0.44446 | -80.6437
0.44814 | -79.5882
0.45181 | -79.717
0.45548 | -80.4418
0.45916 | -81.1626
0.46283 | -79.4281
0.4665 | -79.0405
0.47018 | -79.933
0.47385 | -81.0271
0.47752 | -80.4853
0.4812 | -80.5326
0.48487 | -80.7625
0.48854 | -80.6096
0.49222 | -80.5768
0.49589 | -80.1008
0.49956 | -81.9433
0.50324 | -81.9973
0.50691 | -81.5554
0.51058 | -81.0106
0.51426 | -80.7169
0.51793 | -80.7513
0.5216 | -82.0807
0.52528 | -82.1374
0.52895 | -81.4343
0.53262 | -81.2125
0.5363 | -81.1039
0.53997 | -81.1266
0.54364 | -81.2628
0.54732 | -81.0494
0.55099 | -80.7628
0.55466 | -80.9725
0.55834 | -81.5208
0.56201 | -82.122
0.56568 | -82.5033
0.56936 | -82.4678
0.57303 | -82.1369
0.5767 | -81.9053
0.58038 | -82.154
0.58405 | -82.8558
0.58772 | -83.1458
0.5914 | -82.5761
0.59507 | -82.0091
0.59874 | -82.007
0.60242 | -82.4797
0.60609 | -82.5569
0.60976 | -82.2778
0.61343 | -82.3059
0.61711 | -82.6965
0.62078 | -83.336
0.62445 | -83.5157
0.62813 | -82.8244
0.6318 | -82.1553
0.63548 | -82.0083
0.63915 | -82.1714
0.64282 | -82.5171
0.64649 | -82.9186
0.65017 | -82.9594
0.65384 | -83.4457
0.65751 | -84.1581
0.66119 | -84.2288
0.66486 | -84.1077
0.66853 | -83.8893
0.67221 | -83.6539
0.67588 | -83.668
0.67955 | -83.4257
0.68323 | -83.0439
0.6869 | -82.6385
0.69057 | -82.4329
0.69425 | -83.0582
0.69792 | -84.3696
0.70159 | -85.2367
0.70527 | -85.2407
0.70894 | -84.8752
0.71261 | -84.1673
0.71629 | -83.4688
0.71996 | -83.1987
0.72363 | -83.653
0.72731 | -84.0652
0.73098 | -84.0384
0.73465 | -84.404
0.73833 | -85.14
0.742 | -85.3709
0.74567 | -85.1266
0.74935 | -84.8336
0.75302 | -84.7389
0.75669 | -84.8392
0.76036 | -85.2463
0.76404 | -85.8883
0.76771 | -86.0271
0.77138 | -85.4668
0.77506 | -84.8108
0.77873 | -84.5438
0.78241 | -84.7676
0.78608 | -85.0808
0.78975 | -85.2748
0.79343 | -85.5341
0.7971 | -85.4295
0.80077 | -84.8856
0.80444 | -84.7743
0.80812 | -85.0669
0.81179 | -85.4391
0.81546 | -85.4809
0.81914 | -85.2697
0.82281 | -85.0663
0.82648 | -85.2474
0.83016 | -85.542
0.83383 | -85.5611
0.8375 | -85.6067
0.84118 | -85.6753
0.84485 | -85.8282
0.84852 | -86.3241
0.8522 | -86.5006
0.85587 | -86.3825
0.85954 | -86.2678
0.86322 | -86.1566
0.86689 | -85.966
0.87056 | -85.7701
0.87424 | -85.9453
0.87791 | -86.2088
0.88158 | -86.0654
0.88526 | -85.9035
0.88893 | -86.0544
0.8926 | -86.1589
0.89628 | -86.0451
0.89995 | -86.2486
0.90362 | -86.2179
0.9073 | -86.028
0.91097 | -86.2181
0.91464 | -86.8774
0.91832 | -87.5489
0.92199 | -87.4429
0.92566 | -86.9976
0.92934 | -86.6302
0.93301 | -86.6353
0.93668 | -86.8562
0.94036 | -86.8589
0.94403 | -86.5539
0.9477 | -86.4823
0.95137 | -86.3529
0.95505 | -86.4008
0.95872 | -86.8548
0.9624 | -87.0721
0.96607 | -86.9127
0.96974 | -86.5668
0.97342 | -86.6016
0.97709 | -87.2207
0.98076 | -87.5609
0.98443 | -87.2259
0.98811 | -86.9541
0.99178 | -87.0507
0.99545 | -87.1743
0.99913 | -87.2758
1.0028 | -87.402
1.0065 | -87.2159
1.0101 | -87.0093
1.0138 | -87.0258
1.0175 | -87.2191
1.0212 | -87.3332
1.0248 | -87.2259
1.0285 | -87.5097
1.0322 | -88.3686
1.0359 | -88.7897
1.0395 | -87.9794
1.0432 | -87.1802
1.0469 | -86.9206
1.0506 | -86.747
1.0542 | -86.8142
1.0579 | -87.329
1.0616 | -87.6925
1.0652 | -87.911
1.0689 | -88.0572
1.0726 | -88.2626
1.0763 | -88.5635
1.0799 | -88.7696
1.0836 | -88.7684
1.0873 | -88.7731
1.091 | -88.8189
1.0946 | -88.6635
1.0983 | -88.3592
1.102 | -88.2161
1.1057 | -88.2983
1.1093 | -88.4104
1.113 | -88.4424
1.1167 | -88.4808
1.1203 | -88.6016
1.124 | -88.8292
1.1277 | -89.1002
1.1314 | -89.2219
1.135 | -89.067
1.1387 | -88.7018
1.1424 | -88.3208
1.1461 | -88.1882
1.1497 | -88.3876
1.1534 | -88.6685
1.1571 | -88.6883
1.1608 | -88.5261
1.1644 | -88.4977
1.1681 | -88.6535
1.1718 | -88.8225
1.1754 | -88.8709
1.1791 | -88.8372
1.1828 | -88.7834
1.1865 | -88.7548
1.1901 | -88.7392
1.1938 | -88.6466
1.1975 | -88.6153
1.2012 | -88.8494
1.2048 | -89.1574
1.2085 | -89.215
1.2122 | -89.0722
1.2159 | -88.9447
1.2195 | -88.9259
1.2232 | -89.0086
1.2269 | -89.1251
1.2305 | -89.2543
1.2342 | -89.3746
1.2379 | -89.4014
1.2416 | -89.4202
1.2452 | -89.5631
1.2489 | -89.7172
1.2526 | -89.7543
1.2563 | -89.794
1.2599 | -89.8515
1.2636 | -89.7853
1.2673 | -89.5932
1.2709 | -89.4359
1.2746 | -89.5053
1.2783 | -89.7382
1.282 | -89.8461
1.2856 | -89.7771
1.2893 | -89.5567
1.293 | -89.2331
1.2967 | -89.0435
1.3003 | -89.0697
1.304 | -89.2376
1.3077 | -89.4634
1.3114 | -89.5672
1.315 | -89.4649
1.3187 | -89.3445
1.3224 | -89.3629
1.326 | -89.4797
1.3297 | -89.6575
1.3334 | -89.9361
1.3371 | -90.1977
1.3407 | -90.2238
1.3444 | -89.9349
1.3481 | -89.5657
1.3518 | -89.5206
1.3554 | -89.83
1.3591 | -90.0384
1.3628 | -89.9414
1.3665 | -89.9337
1.3701 | -90.1855
1.3738 | -90.2675
1.3775 | -90.0117
1.3811 | -89.7727
1.3848 | -89.6838
1.3885 | -89.7293
1.3922 | -89.9603
1.3958 | -90.2844
1.3995 | -90.4304
1.4032 | -90.27
1.4069 | -90.0875
1.4105 | -90.0616
1.4142 | -90.0855
1.4179 | -90.1091
1.4216 | -90.1997
1.4252 | -90.3774
1.4289 | -90.6304
1.4326 | -90.8879
1.4362 | -91.1273
1.4399 | -91.4021
1.4436 | -91.6001
1.4473 | -91.457
1.4509 | -90.9551
1.4546 | -90.4072
1.4583 | -90.1299
1.462 | -90.2275
1.4656 | -90.5627
1.4693 | -90.7004
1.473 | -90.3801
1.4767 | -89.9788
1.4803 | -89.8374
1.484 | -89.972
1.4877 | -90.2698
1.4913 | -90.5971
1.495 | -90.8016
1.4987 | -90.8083
1.5024 | -90.8086
1.506 | -90.984
1.5097 | -91.2044
1.5134 | -91.18
1.5171 | -90.9491
1.5207 | -90.7975
1.5244 | -90.7998
1.5281 | -90.8244
1.5318 | -90.8202
1.5354 | -90.8368
1.5391 | -90.9348
1.5428 | -91.1303
1.5464 | -91.3102
1.5501 | -91.3333
1.5538 | -91.2209
1.5575 | -91.0717
1.5611 | -90.9731
1.5648 | -90.9133
1.5685 | -90.9038
1.5722 | -91.0735
1.5758 | -91.3379
1.5795 | -91.4513
1.5832 | -91.4552
1.5869 | -91.4924
1.5905 | -91.4891
1.5942 | -91.4223
1.5979 | -91.3117
1.6015 | -91.1988
1.6052 | -91.0959
1.6089 | -90.9881
1.6126 | -90.8813
1.6162 | -90.8661
1.6199 | -90.9856
1.6236 | -91.1532
1.6273 | -91.2325
1.6309 | -91.1977
1.6346 | -91.1362
1.6383 | -91.1429
1.6419 | -91.2262
1.6456 | -91.3413
1.6493 | -91.4404
1.653 | -91.5324
1.6566 | -91.6644
1.6603 | -91.8795
1.664 | -92.1542
1.6677 | -92.3857
1.6713 | -92.4579
1.675 | -92.3799
1.6787 | -92.2432
1.6824 | -92.1255
1.686 | -92.0686
1.6897 | -92.0495
1.6934 | -92.0246
1.697 | -91.9744
1.7007 | -91.9131
1.7044 | -91.8672
1.7081 | -91.8247
1.7117 | -91.789
1.7154 | -91.8214
1.7191 | -91.9935
1.7228 | -92.2954
1.7264 | -92.5853
1.7301 | -92.7457
1.7338 | -92.8173
1.7375 | -92.876
1.7411 | -92.8335
1.7448 | -92.5149
1.7485 | -91.9305
1.7521 | -91.332
1.7558 | -90.9903
1.7595 | -91.0265
1.7632 | -91.3501
1.7668 | -91.7284
1.7705 | -91.9714
1.7742 | -92.0824
1.7779 | -92.1156
1.7815 | -92.0916
1.7852 | -92.0454
1.7889 | -92.0521
1.7926 | -92.1868
1.7962 | -92.4221
1.7999 | -92.6156
1.8036 | -92.6151
1.8072 | -92.4031
1.8109 | -92.0941
1.8146 | -91.8604
1.8183 | -91.8007
1.8219 | -91.924
1.8256 | -92.1562
1.8293 | -92.383
1.833 | -92.5091
1.8366 | -92.5408
1.8403 | -92.5488
1.844 | -92.5978
1.8477 | -92.6802
1.8513 | -92.8091
1.855 | -92.9877
1.8587 | -93.1819
1.8623 | -93.3392
1.866 | -93.4293
1.8697 | -93.4566
1.8734 | -93.4139
1.877 | -93.2686
1.8807 | -93.069
1.8844 | -92.8935
1.8881 | -92.7906
1.8917 | -92.7586
1.8954 | -92.7952
1.8991 | -92.9111
1.9027 | -93.0991
1.9064 | -93.2917
1.9101 | -93.373
1.9138 | -93.252
1.9174 | -92.9554
1.9211 | -92.5977
1.9248 | -92.3135
1.9285 | -92.2114
1.9321 | -92.3435
1.9358 | -92.7
1.9395 | -93.1805
1.9432 | -93.6054
1.9468 | -93.8151
1.9505 | -93.7664
1.9542 | -93.5198
1.9578 | -93.2185
1.9615 | -92.9869
1.9652 | -92.8591
1.9689 | -92.7725
1.9725 | -92.658
1.9762 | -92.5569
1.9799 | -92.5727
1.9836 | -92.7605
1.9872 | -93.0658
1.9909 | -93.3782
1.9946 | -93.6252
1.9983 | -93.77
2.0019 | -93.7993
2.0056 | -93.7698
2.0093 | -93.7063
2.0129 | -93.5827
2.0166 | -93.3717
2.0203 | -93.1033
2.024 | -92.8195
2.0276 | -92.6092
2.0313 | -92.5612
2.035 | -92.7034
2.0387 | -92.9792
2.0423 | -93.3066
2.046 | -93.5705
2.0497 | -93.6443
2.0534 | -93.4771
2.057 | -93.226
2.0607 | -93.076
2.0644 | -93.1061
2.068 | -93.271
2.0717 | -93.4957
2.0754 | -93.6881
2.0791 | -93.7897
2.0827 | -93.8571
2.0864 | -93.9804
2.0901 | -94.1215
2.0938 | -94.1764
2.0974 | -94.1306
2.1011 | -94.0478
2.1048 | -93.9595
2.1085 | -93.869
2.1121 | -93.7943
2.1158 | -93.7459
2.1195 | -93.7483
2.1231 | -93.7149
2.1268 | -93.648
2.1305 | -93.5303
2.1342 | -93.3922
2.1378 | -93.2812
2.1415 | -93.251
2.1452 | -93.3236
2.1489 | -93.4788
2.1525 | -93.6674
2.1562 | -93.825
2.1599 | -93.936
2.1636 | -94.0063
2.1672 | -94.0539
2.1709 | -94.1096
2.1746 | -94.178
2.1782 | -94.2606
2.1819 | -94.3333
2.1856 | -94.354
2.1893 | -94.3234
2.1929 | -94.2603
2.1966 | -94.1636
2.2003 | -94.0485
2.204 | -93.9636
2.2076 | -93.9302
2.2113 | -93.9487
2.215 | -93.9549
2.2187 | -93.8722
2.2223 | -93.6996
2.226 | -93.5135
2.2297 | -93.3992
2.2333 | -93.421
2.237 | -93.5968
2.2407 | -93.872
2.2444 | -94.1576
2.248 | -94.3757
2.2517 | -94.453
2.2554 | -94.4111
2.2591 | -94.3106
2.2627 | -94.1903
2.2664 | -94.0701
2.2701 | -93.9383
2.2738 | -93.7842
2.2774 | -93.6312
2.2811 | -93.5289
2.2848 | -93.5156
2.2884 | -93.5987
2.2921 | -93.7447
2.2958 | -93.8976
2.2995 | -93.9899
2.3031 | -93.9966
2.3068 | -93.9702
2.3105 | -93.9771
2.3142 | -94.027
2.3178 | -94.0782
2.3215 | -94.1222
2.3252 | -94.197
2.3288 | -94.337
2.3325 | -94.5115
2.3362 | -94.6268
2.3399 | -94.6203
2.3435 | -94.4902
2.3472 | -94.2983
2.3509 | -94.1234
2.3546 | -94.0166
2.3582 | -93.9746
2.3619 | -93.9793
2.3656 | -93.99
2.3693 | -93.98
2.3729 | -93.9362
2.3766 | -93.8943
2.3803 | -93.9117
2.3839 | -94.0274
2.3876 | -94.2216
2.3913 | -94.4424
2.395 | -94.6264
2.3986 | -94.7159
2.4023 | -94.6718
2.406 | -94.5229
2.4097 | -94.3352
2.4133 | -94.1914
2.417 | -94.1428
2.4207 | -94.203
2.4244 | -94.3452
2.428 | -94.5056
2.4317 | -94.6027
2.4354 | -94.5978
2.439 | -94.5447
2.4427 | -94.5076
2.4464 | -94.5
2.4501 | -94.5193
2.4537 | -94.581
2.4574 | -94.6996
2.4611 | -94.8803
2.4648 | -95.0838
2.4684 | -95.2481
2.4721 | -95.3308
2.4758 | -95.3316
2.4795 | -95.2454
2.4831 | -95.0844
2.4868 | -94.8999
2.4905 | -94.7299
2.4941 | -94.6273
2.4978 | -94.622
2.5015 | -94.6813
2.5052 | -94.7568
2.5088 | -94.8007
2.5125 | -94.8067
2.5162 | -94.7864
2.5199 | -94.7643
2.5235 | -94.7509
2.5272 | -94.7483
2.5309 | -94.7775
2.5346 | -94.8522
2.5382 | -94.9781
2.5419 | -95.1495
2.5456 | -95.3002
2.5492 | -95.3694
2.5529 | -95.3265
2.5566 | -95.2105
2.5603 | -95.0854
2.5639 | -94.97
2.5676 | -94.8469
2.5713 | -94.7352
2.575 | -94.684
2.5786 | -94.7205
2.5823 | -94.832
2.586 | -95.0013
2.5897 | -95.2069
2.5933 | -95.42
2.597 | -95.6113
2.6007 | -95.7647
2.6043 | -95.8876
2.608 | -95.9981
2.6117 | -96.0778
2.6154 | -96.0758
2.619 | -95.9712
2.6227 | -95.7876
2.6264 | -95.5995
2.6301 | -95.482
2.6337 | -95.4654
2.6374 | -95.5302
2.6411 | -95.6214
2.6448 | -95.7105
2.6484 | -95.7015
2.6521 | -95.6922
2.6558 | -95.624
2.6594 | -95.5141
2.6631 | -95.3889
2.6668 | -95.2768
2.6705 | -95.1957
2.6741 | -95.1581
2.6778 | -95.1728
2.6815 | -95.229
2.6852 | -95.3291
2.6888 | -95.4393
2.6925 | -95.5154
2.6962 | -95.5359
2.6998 | -95.5089
2.7035 | -95.4929
2.7072 | -95.5267
2.7109 | -95.6248
2.7145 | -95.7672
2.7182 | -95.9215
2.7219 | -96.0453
2.7256 | -96.1056
2.7292 | -96.1063
2.7329 | -96.0402
2.7366 | -95.9406
2.7403 | -95.8401
2.7439 | -95.7572
2.7476 | -95.7039
2.7513 | -95.6941
2.7549 | -95.7386
2.7586 | -95.8209
2.7623 | -95.9105
2.766 | -95.9725
2.7696 | -95.9954
2.7733 | -95.985
2.777 | -95.9411
2.7807 | -95.8857
2.7843 | -95.7978
2.788 | -95.6907
2.7917 | -95.586
2.7954 | -95.4944
2.799 | -95.4158
2.8027 | -95.3591
2.8064 | -95.3651
2.81 | -95.4336
2.8137 | -95.5587
2.8174 | -95.7187
2.8211 | -95.8751
2.8247 | -96.0081
2.8284 | -96.0982
2.8321 | -96.1591
2.8358 | -96.1848
2.8394 | -96.1764
2.8431 | -96.1717
2.8468 | -96.1887
2.8505 | -96.2206
2.8541 | -96.2673
2.8578 | -96.3578
2.8615 | -96.4633
2.8651 | -96.5659
2.8688 | -96.6572
2.8725 | -96.7195
2.8762 | -96.7645
2.8798 | -96.8025
2.8835 | -96.8346
2.8872 | -96.8624
2.8909 | -96.8959
2.8945 | -96.9418
2.8982 | -96.9893
2.9019 | -97.0078
2.9056 | -96.9544
2.9092 | -96.8488
2.9129 | -96.7169
2.9166 | -96.6215
2.9202 | -96.5885
2.9239 | -96.5989
2.9276 | -96.6192
2.9313 | -96.6344
2.9349 | -96.6523
2.9386 | -96.6701
2.9423 | -96.7093
2.946 | -96.762
2.9496 | -96.8177
2.9533 | -96.8698
2.957 | -96.909
2.9606 | -96.9082
2.9643 | -96.8375
2.968 | -96.716
2.9717 | -96.5765
2.9753 | -96.4652
2.979 | -96.3948
2.9827 | -96.3612
2.9864 | -96.3714
2.99 | -96.4101
2.9937 | -96.4757
2.9974 | -96.553
3.0011 | -96.6502
3.0047 | -96.7478
3.0084 | -96.8498
3.0121 | -96.9405
3.0157 | -96.9891
3.0194 | -96.9704
3.0231 | -96.8709
3.0268 | -96.7429
3.0304 | -96.6479
3.0341 | -96.6187
3.0378 | -96.682
3.0415 | -96.7967
3.0451 | -96.9138
3.0488 | -96.9879
3.0525 | -97.0153
3.0562 | -97.0109
3.0598 | -97.0077
3.0635 | -97.0193
3.0672 | -97.0572
3.0708 | -97.1127
3.0745 | -97.1567
3.0782 | -97.1747
3.0819 | -97.1557
3.0855 | -97.1089
3.0892 | -97.0709
3.0929 | -97.0623
3.0966 | -97.0911
3.1002 | -97.148
3.1039 | -97.2404
3.1076 | -97.3347
3.1113 | -97.4279
3.1149 | -97.5028
3.1186 | -97.5539
3.1223 | -97.5863
3.1259 | -97.6008
3.1296 | -97.6226
3.1333 | -97.6251
3.137 | -97.5941
3.1406 | -97.5289
3.1443 | -97.4246
3.148 | -97.312
3.1517 | -97.1885
3.1553 | -97.0729
3.159 | -96.9577
3.1627 | -96.8561
3.1664 | -96.7684
3.17 | -96.7032
3.1737 | -96.6783
3.1774 | -96.757
3.181 | -96.8089
3.1847 | -96.895
3.1884 | -96.9853
3.1921 | -97.0532
3.1957 | -97.0939
3.1994 | -97.0943
3.2031 | -97.063
3.2068 | -97.0193
3.2104 | -96.9758
3.2141 | -96.9474
3.2178 | -96.9423
3.2215 | -96.9566
3.2251 | -96.9902
3.2288 | -97.0308
3.2325 | -97.0685
3.2361 | -97.1133
3.2398 | -97.1481
3.2435 | -97.2055
3.2472 | -97.2945
3.2508 | -97.41
3.2545 | -97.5433
3.2582 | -97.6777
3.2619 | -97.7845
3.2655 | -97.8496
3.2692 | -97.8533
3.2729 | -97.8155
3.2766 | -97.7468
3.2802 | -97.6539
3.2839 | -97.5443
3.2876 | -97.4209
3.2912 | -97.2892
3.2949 | -97.1628
3.2986 | -97.0461
3.3023 | -96.9684
3.3059 | -96.9326
3.3096 | -96.9414
3.3133 | -96.9909
3.317 | -97.0781
3.3206 | -97.2033
3.3243 | -97.3671
3.328 | -97.5546
3.3316 | -97.7582
3.3353 | -97.9668
3.339 | -98.1522
3.3427 | -98.2891
3.3463 | -98.3566
3.35 | -98.3621
3.3537 | -98.323
3.3574 | -98.2563
3.361 | -98.1979
3.3647 | -98.1524
3.3684 | -98.1248
3.3721 | -98.109
3.3757 | -98.097
3.3794 | -98.0888
3.3831 | -98.0684
3.3867 | -98.0347
3.3904 | -97.9941
3.3941 | -97.9602
3.3978 | -97.9305
3.4014 | -97.9153
3.4051 | -97.9133
3.4088 | -97.9148
3.4125 | -97.9159
3.4161 | -97.901
3.4198 | -97.8658
3.4235 | -97.8067
3.4272 | -97.7332
3.4308 | -97.6622
3.4345 | -97.5972
3.4382 | -97.5504
3.4418 | -97.5202
3.4455 | -97.5151
3.4492 | -97.5268
3.4529 | -97.5514
3.4565 | -97.5814
3.4602 | -97.6271
3.4639 | -97.6766
3.4676 | -97.7198
3.4712 | -97.7597
3.4749 | -97.7921
3.4786 | -97.8197
3.4823 | -97.8523
3.4859 | -97.891
3.4896 | -97.9334
3.4933 | -97.975
3.4969 | -98.0069
3.5006 | -98.0192
3.5043 | -98.0197
3.508 | -98.0066
3.5116 | -97.9868
3.5153 | -97.9567
3.519 | -97.94
3.5227 | -97.9448
3.5263 | -97.9649
3.53 | -98.0099
3.5337 | -98.0785
3.5374 | -98.1585
3.541 | -98.2529
3.5447 | -98.3516
3.5484 | -98.4313
3.552 | -98.49
3.5557 | -98.53
3.5594 | -98.5503
3.5631 | -98.5485
3.5667 | -98.5199
3.5704 | -98.467
3.5741 | -98.3705
3.5778 | -98.2335
3.5814 | -98.0825
3.5851 | -97.9475
3.5888 | -97.8674
3.5925 | -97.8718
3.5961 | -97.9416
3.5998 | -98.0562
3.6035 | -98.1728
3.6071 | -98.2541
3.6108 | -98.2894
3.6145 | -98.3006
3.6182 | -98.3049
3.6218 | -98.3219
3.6255 | -98.3617
3.6292 | -98.4392
3.6329 | -98.5463
3.6365 | -98.6567
3.6402 | -98.7497
3.6439 | -98.825
3.6475 | -98.8626
3.6512 | -98.8663
3.6549 | -98.8302
3.6586 | -98.7625
3.6622 | -98.6785
3.6659 | -98.5909
3.6696 | -98.4997
3.6733 | -98.4003
3.6769 | -98.2914
3.6806 | -98.1637
3.6843 | -98.0139
3.688 | -97.8713
3.6916 | -97.765
3.6953 | -97.7119
3.699 | -97.7255
3.7026 | -97.8074
3.7063 | -97.9739
3.71 | -98.1
3.7137 | -98.2335
3.7173 | -98.3492
3.721 | -98.4365
3.7247 | -98.4837
3.7284 | -98.4931
3.732 | -98.4885
3.7357 | -98.4699
3.7394 | -98.4446
3.7431 | -98.4161
3.7467 | -98.3732
3.7504 | -98.3272
3.7541 | -98.2927
3.7577 | -98.2719
3.7614 | -98.2819
3.7651 | -98.3229
3.7688 | -98.3828
3.7724 | -98.4592
3.7761 | -98.5375
3.7798 | -98.6193
3.7835 | -98.699
3.7871 | -98.7741
3.7908 | -98.8311
3.7945 | -98.8598
3.7982 | -98.8748
3.8018 | -98.8898
3.8055 | -98.9126
3.8092 | -98.9483
3.8128 | -98.9855
3.8165 | -99.0197
3.8202 | -99.0508
3.8239 | -99.0686
3.8275 | -99.0736
3.8312 | -99.0815
3.8349 | -99.081
3.8386 | -99.0827
3.8422 | -99.0877
3.8459 | -99.0921
3.8496 | -99.0923
3.8533 | -99.0775
3.8569 | -99.038
3.8606 | -98.966
3.8643 | -98.8666
3.8679 | -98.7676
3.8716 | -98.6838
3.8753 | -98.6282
3.879 | -98.6053
3.8826 | -98.6173
3.8863 | -98.6582
3.89 | -98.7172
3.8937 | -98.7861
3.8973 | -98.8682
3.901 | -98.9536
3.9047 | -99.0352
3.9084 | -99.1138
3.912 | -99.1854
3.9157 | -99.2431
3.9194 | -99.2792
3.923 | -99.3007
3.9267 | -99.318
3.9304 | -99.316
3.9341 | -99.3134
3.9377 | -99.3121
3.9414 | -99.3071
3.9451 | -99.3005
3.9488 | -99.2902
3.9524 | -99.2729
3.9561 | -99.2604
3.9598 | -99.2503
3.9635 | -99.2537
3.9671 | -99.2577
3.9708 | -99.2557
3.9745 | -99.257
3.9781 | -99.2602
3.9818 | -99.2634
3.9855 | -99.2606
3.9892 | -99.245
3.9928 | -99.2204
3.9965 | -99.1777
4.0002 | -99.1249
4.0039 | -99.0652
4.0075 | -98.9975
4.0112 | -98.9404
4.0149 | -98.9098
4.0186 | -98.9089
4.0222 | -98.9452
4.0259 | -99.019
4.0296 | -99.1327
4.0332 | -99.2685
4.0369 | -99.4053
4.0406 | -99.5307
4.0443 | -99.6395
4.0479 | -99.721
4.0516 | -99.7639
4.0553 | -99.767
4.059 | -99.7277
4.0626 | -99.6582
4.0663 | -99.587
4.07 | -99.531
4.0736 | -99.4988
4.0773 | -99.4962
4.081 | -99.5291
4.0847 | -99.5832
4.0883 | -99.6485
4.092 | -99.7045
4.0957 | -99.7455
4.0994 | -99.7695
4.103 | -99.7614
4.1067 | -99.7209
4.1104 | -99.6471
4.1141 | -99.5629
4.1177 | -99.4848
4.1214 | -99.432
4.1251 | -99.4183
4.1287 | -99.4353
4.1324 | -99.4856
4.1361 | -99.5566
4.1398 | -99.6341
4.1434 | -99.7035
4.1471 | -99.7645
4.1508 | -99.8147
4.1545 | -99.8532
4.1581 | -99.8766
4.1618 | -99.8965
4.1655 | -99.9229
4.1692 | -99.9459
4.1728 | -99.9739
4.1765 | -99.9914
4.1802 | -99.9927
4.1838 | -99.9732
4.1875 | -99.9266
4.1912 | -99.8678
4.1949 | -99.7898
4.1985 | -99.6929
4.2022 | -99.6014
4.2059 | -99.5221
4.2096 | -99.4681
4.2132 | -99.4441
4.2169 | -99.4666
4.2206 | -99.5236
4.2243 | -99.6042
4.2279 | -99.7008
4.2316 | -99.8032
4.2353 | -99.8936
4.2389 | -99.9475
4.2426 | -100.0135
4.2463 | -100.0657
4.25 | -100.1083
4.2536 | -100.1357
4.2573 | -100.1531
4.261 | -100.1725
4.2647 | -100.1894
4.2683 | -100.1938
4.272 | -100.1863
4.2757 | -100.1629
4.2794 | -100.1235
4.283 | -100.0646
4.2867 | -99.9994
4.2904 | -99.9309
4.294 | -99.8744
4.2977 | -99.829
4.3014 | -99.7962
4.3051 | -99.7762
4.3087 | -99.7608
4.3124 | -99.7527
4.3161 | -99.7482
4.3198 | -99.7446
4.3234 | -99.7411
4.3271 | -99.7271
4.3308 | -99.7195
4.3345 | -99.7229
4.3381 | -99.732
4.3418 | -99.749
4.3455 | -99.7725
4.3491 | -99.8035
4.3528 | -99.8406
4.3565 | -99.8817
4.3602 | -99.9299
4.3638 | -99.9818
4.3675 | -100.0434
4.3712 | -100.1171
4.3749 | -100.1911
4.3785 | -100.2728
4.3822 | -100.3536
4.3859 | -100.4228
4.3895 | -100.4869
4.3932 | -100.5306
4.3969 | -100.5527
4.4006 | -100.5512
4.4042 | -100.5156
4.4079 | -100.4613
4.4116 | -100.3872
4.4153 | -100.2799
4.4189 | -100.155
4.4226 | -100.0259
4.4263 | -99.9044
4.43 | -99.7979
4.4336 | -99.7217
4.4373 | -99.6866
4.441 | -99.697
4.4446 | -99.7496
4.4483 | -99.8344
4.452 | -99.9385
4.4557 | -100.046
4.4593 | -100.1443
4.463 | -100.2266
4.4667 | -100.2852
4.4704 | -100.3129
4.474 | -100.3197
4.4777 | -100.2889
4.4814 | -100.2334
4.4851 | -100.1712
4.4887 | -100.0938
4.4924 | -100.0103
4.4961 | -99.935
4.4997 | -99.8853
4.5034 | -99.8484
4.5071 | -99.8308
4.5108 | -99.8352
4.5144 | -99.8692
4.5181 | -99.9276
4.5218 | -100.0098
4.5255 | -100.1073
4.5291 | -100.2083
4.5328 | -100.3122
4.5365 | -100.3999
4.5402 | -100.4605
4.5438 | -100.4876
4.5475 | -100.4864
4.5512 | -100.4435
4.5548 | -100.3711
4.5585 | -100.2949
4.5622 | -100.2228
4.5659 | -100.153
4.5695 | -100.1005
4.5732 | -100.0698
4.5769 | -100.0457
4.5806 | -100.0307
4.5842 | -100.021
4.5879 | -100.0139
4.5916 | -100.0122
4.5953 | -100.0154
4.5989 | -100.0302
4.6026 | -100.0624
4.6063 | -100.1099
4.6099 | -100.1756
4.6136 | -100.2584
4.6173 | -100.3527
4.621 | -100.4531
4.6246 | -100.5417
4.6283 | -100.6072
4.632 | -100.6476
4.6357 | -100.668
4.6393 | -100.6599
4.643 | -100.6288
4.6467 | -100.5803
4.6504 | -100.5195
4.654 | -100.455
4.6577 | -100.4004
4.6614 | -100.3667
4.665 | -100.3544
4.6687 | -100.3517
4.6724 | -100.36
4.6761 | -100.3809
4.6797 | -100.4009
4.6834 | -100.4165
4.6871 | -100.4148
4.6908 | -100.4082
4.6944 | -100.3928
4.6981 | -100.3707
4.7018 | -100.3489
4.7054 | -100.3235
4.7091 | -100.3049
4.7128 | -100.2963
4.7165 | -100.2956
4.7201 | -100.3025
4.7238 | -100.3279
4.7275 | -100.3695
4.7312 | -100.4084
4.7348 | -100.454
4.7385 | -100.5085
4.7422 | -100.5463
4.7459 | -100.5748
4.7495 | -100.6024
4.7532 | -100.6242
4.7569 | -100.6526
4.7605 | -100.6829
4.7642 | -100.7193
4.7679 | -100.7601
4.7716 | -100.7919
4.7752 | -100.8203
4.7789 | -100.8463
4.7826 | -100.8669
4.7863 | -100.8867
4.7899 | -100.8996
4.7936 | -100.8965
4.7973 | -100.8742
4.801 | -100.8409
4.8046 | -100.8064
4.8083 | -100.7695
4.812 | -100.7349
4.8156 | -100.7224
4.8193 | -100.7212
4.823 | -100.7252
4.8267 | -100.7378
4.8303 | -100.7468
4.834 | -100.7504
4.8377 | -100.7567
4.8414 | -100.7645
4.845 | -100.7731
4.8487 | -100.7869
4.8524 | -100.8046
4.8561 | -100.835
4.8597 | -100.8758
4.8634 | -100.9196
4.8671 | -100.966
4.8707 | -101.0169
4.8744 | -101.0643
4.8781 | -101.1037
4.8818 | -101.1332
4.8854 | -101.1553
4.8891 | -101.1688
4.8928 | -101.1687
4.8965 | -101.1612
4.9001 | -101.1491
4.9038 | -101.1277
4.9075 | -101.0961
4.9112 | -101.0601
4.9148 | -101.0145
4.9185 | -100.9639
4.9222 | -100.9098
4.9258 | -100.8581
4.9295 | -100.8167
4.9332 | -100.7799
4.9369 | -100.7561
4.9405 | -100.7546
4.9442 | -100.7702
4.9479 | -100.806
4.9516 | -100.8606
4.9552 | -100.9293
4.9589 | -101.0066
4.9626 | -101.0826
4.9663 | -101.1522
4.9699 | -101.2123
4.9736 | -101.2553
4.9773 | -101.2821
4.9809 | -101.292
4.9846 | -101.2847
4.9883 | -101.2731
4.992 | -101.2656
4.9956 | -101.2615
4.9993 | -101.2676
5.003 | -101.2665
5.0067 | -101.2624
5.0103 | -101.2589
5.014 | -101.2488
5.0177 | -101.23
5.0214 | -101.2105
5.025 | -101.1852
5.0287 | -101.1627
5.0324 | -101.1445
5.036 | -101.1299
5.0397 | -101.1148
5.0434 | -101.1077
5.0471 | -101.1054
5.0507 | -101.1013
5.0544 | -101.1041
5.0581 | -101.1106
5.0618 | -101.1255
5.0654 | -101.1457
5.0691 | -101.1757
5.0728 | -101.2192
5.0765 | -101.257
5.0801 | -101.3011
5.0838 | -101.3363
5.0875 | -101.3528
5.0911 | -101.3475
5.0948 | -101.3248
5.0985 | -101.2892
5.1022 | -101.25
5.1058 | -101.2179
5.1095 | -101.2121
5.1132 | -101.2416
5.1169 | -101.305
5.1205 | -101.3826
5.1242 | -101.473
5.1279 | -101.5653
5.1315 | -101.6313
5.1352 | -101.6631
5.1389 | -101.6457
5.1426 | -101.5758
5.1462 | -101.4542
5.1499 | -101.2988
5.1536 | -101.1337
5.1573 | -100.9618
5.1609 | -100.8039
5.1646 | -100.6727
5.1683 | -100.5814
5.172 | -100.5299
5.1756 | -100.5196
5.1793 | -100.5446
5.183 | -100.5916
5.1866 | -100.6506
5.1903 | -100.724
5.194 | -100.8049
5.1977 | -100.8937
5.2013 | -100.9798
5.205 | -101.0642
5.2087 | -101.1497
5.2124 | -101.2364
5.216 | -101.3252
5.2197 | -101.4182
5.2234 | -101.5139
5.2271 | -101.6054
5.2307 | -101.695
5.2344 | -101.7706
5.2381 | -101.8237
5.2417 | -101.8597
5.2454 | -101.8568
5.2491 | -101.8149
5.2528 | -101.756
5.2564 | -101.6874
5.2601 | -101.6151
5.2638 | -101.5442
5.2675 | -101.4735
5.2711 | -101.4213
5.2748 | -101.3912
5.2785 | -101.377
5.2822 | -101.3748
5.2858 | -101.3956
5.2895 | -101.431
5.2932 | -101.4805
5.2968 | -101.5357
5.3005 | -101.5841
5.3042 | -101.6314
5.3079 | -101.6667
5.3115 | -101.6832
5.3152 | -101.6991
5.3189 | -101.7114
5.3226 | -101.7114
5.3262 | -101.7012
5.3299 | -101.6911
5.3336 | -101.6733
5.3373 | -101.6495
5.3409 | -101.627
5.3446 | -101.6062
5.3483 | -101.5801
5.3519 | -101.5472
5.3556 | -101.5087
5.3593 | -101.4696
5.363 | -101.4289
5.3666 | -101.3781
5.3703 | -101.3231
5.374 | -101.265
5.3777 | -101.2061
5.3813 | -101.1616
5.385 | -101.1179
5.3887 | -101.089
5.3924 | -101.0809
5.396 | -101.0943
5.3997 | -101.1252
5.4034 | -101.1774
5.407 | -101.2519
5.4107 | -101.3471
5.4144 | -101.4482
5.4181 | -101.5478
5.4217 | -101.653
5.4254 | -101.7578
5.4291 | -101.8559
5.4328 | -101.9435
5.4364 | -102.0124
5.4401 | -102.0579
5.4438 | -102.0771
5.4474 | -102.0651
5.4511 | -102.0203
5.4548 | -101.9573
5.4585 | -101.8807
5.4621 | -101.7996
5.4658 | -101.7131
5.4695 | -101.6342
5.4732 | -101.5619
5.4768 | -101.493
5.4805 | -101.4378
5.4842 | -101.3991
5.4879 | -101.3762
5.4915 | -101.3753
5.4952 | -101.3944
5.4989 | -101.4305
5.5025 | -101.4766
5.5062 | -101.5247
5.5099 | -101.569
5.5136 | -101.6088
5.5172 | -101.6481
5.5209 | -101.67
5.5246 | -101.6798
5.5283 | -101.6844
5.5319 | -101.6883
5.5356 | -101.6922
5.5393 | -101.6907
5.543 | -101.693
5.5466 | -101.7045
5.5503 | -101.7261
5.554 | -101.7476
5.5576 | -101.7665
5.5613 | -101.7838
5.565 | -101.7912
5.5687 | -101.7837
5.5723 | -101.7716
5.576 | -101.7468
5.5797 | -101.7176
5.5834 | -101.69
5.587 | -101.6747
5.5907 | -101.6701
5.5944 | -101.6774
5.5981 | -101.7026
5.6017 | -101.7432
5.6054 | -101.8034
5.6091 | -101.8715
5.6127 | -101.9427
5.6164 | -102.0128
5.6201 | -102.0752
5.6238 | -102.1254
5.6274 | -102.1544
5.6311 | -102.1679
5.6348 | -102.1538
5.6385 | -102.1202
5.6421 | -102.0592
5.6458 | -101.971
5.6495 | -101.8655
5.6532 | -101.7575
5.6568 | -101.6444
5.6605 | -101.5476
5.6642 | -101.4681
5.6678 | -101.406
5.6715 | -101.3652
5.6752 | -101.348
5.6789 | -101.3504
5.6825 | -101.3693
5.6862 | -101.402
5.6899 | -101.4484
5.6936 | -101.5084
5.6972 | -101.579
5.7009 | -101.6568
5.7046 | -101.7431
5.7083 | -101.828
5.7119 | -101.9032
5.7156 | -101.9624
5.7193 | -102.0188
5.7229 | -102.0531
5.7266 | -102.0681
5.7303 | -102.0656
5.734 | -102.0409
5.7376 | -102.0053
5.7413 | -101.9585
5.745 | -101.9017
5.7487 | -101.8446
5.7523 | -101.7905
5.756 | -101.734
5.7597 | -101.6769
5.7633 | -101.6242
5.767 | -101.5781
5.7707 | -101.5346
5.7744 | -101.503
5.778 | -101.4802
5.7817 | -101.4692
5.7854 | -101.4811
5.7891 | -101.51
5.7927 | -101.556
5.7964 | -101.6099
5.8001 | -101.6692
5.8038 | -101.7307
5.8074 | -101.7981
5.8111 | -101.862
5.8148 | -101.9159
5.8184 | -101.9625
5.8221 | -101.986
5.8258 | -101.9963
5.8295 | -102.0229
5.8331 | -102.0472
5.8368 | -102.0685
5.8405 | -102.0935
5.8442 | -102.1238
5.8478 | -102.1573
5.8515 | -102.19
5.8552 | -102.2218
5.8589 | -102.2615
5.8625 | -102.2926
5.8662 | -102.3135
5.8699 | -102.3308
5.8735 | -102.3379
5.8772 | -102.3329
5.8809 | -102.3212
5.8846 | -102.3102
5.8882 | -102.2956
5.8919 | -102.2727
5.8956 | -102.2451
5.8993 | -102.2206
5.9029 | -102.1984
5.9066 | -102.1817
5.9103 | -102.1757
5.914 | -102.1708
5.9176 | -102.175
5.9213 | -102.1833
5.925 | -102.1979
5.9286 | -102.2178
5.9323 | -102.2389
5.936 | -102.2619
5.9397 | -102.2831
5.9433 | -102.3051
5.947 | -102.3272
5.9507 | -102.3394
5.9544 | -102.3546
5.958 | -102.3762
5.9617 | -102.3803
5.9654 | -102.3739
5.9691 | -102.3705
5.9727 | -102.3579
5.9764 | -102.3352
5.9801 | -102.3062
5.9837 | -102.2788
5.9874 | -102.2517
5.9911 | -102.2247
5.9948 | -102.2002
5.9984 | -102.1737
6.0021 | -102.1518
6.0058 | -102.1331
6.0095 | -102.1113
6.0131 | -102.092
6.0168 | -102.0707
6.0205 | -102.0437
6.0242 | -102.0193
6.0278 | -101.9947
6.0315 | -101.9709
6.0352 | -101.9432
6.0388 | -101.9116
6.0425 | -101.8821
6.0462 | -101.8543
6.0499 | -101.833
6.0535 | -101.824
6.0572 | -101.8221
6.0609 | -101.8249
6.0646 | -101.8392
6.0682 | -101.8649
6.0719 | -101.9037
6.0756 | -101.9483
6.0793 | -101.9984
6.0829 | -102.0521
6.0866 | -102.1116
6.0903 | -102.179
6.0939 | -102.2455
6.0976 | -102.3005
6.1013 | -102.3493
6.105 | -102.3894
6.1086 | -102.4204
6.1123 | -102.4445
6.116 | -102.455
6.1197 | -102.4637
6.1233 | -102.4676
6.127 | -102.4785
6.1307 | -102.4915
6.1344 | -102.5085
6.138 | -102.5279
6.1417 | -102.5416
6.1454 | -102.5538
6.149 | -102.5483
6.1527 | -102.523
6.1564 | -102.4815
6.1601 | -102.434
6.1637 | -102.3638
6.1674 | -102.2837
6.1711 | -102.2019
6.1748 | -102.1237
6.1784 | -102.0481
6.1821 | -101.9764
6.1858 | -101.9153
6.1894 | -101.8789
6.1931 | -101.8638
6.1968 | -101.8678
6.2005 | -101.8998
6.2041 | -101.954
6.2078 | -102.0322
6.2115 | -102.1272
6.2152 | -102.2339
6.2188 | -102.3372
6.2225 | -102.429
6.2262 | -102.5091
6.2299 | -102.5674
6.2335 | -102.6017
6.2372 | -102.6172
6.2409 | -102.621
6.2445 | -102.6063
6.2482 | -102.5755
6.2519 | -102.5327
6.2556 | -102.5004
6.2592 | -102.4786
6.2629 | -102.4616
6.2666 | -102.449
6.2703 | -102.4438
6.2739 | -102.4451
6.2776 | -102.4521
6.2813 | -102.452
6.285 | -102.4438
6.2886 | -102.437
6.2923 | -102.424
6.296 | -102.4178
6.2996 | -102.4103
6.3033 | -102.4102
6.307 | -102.4238
6.3107 | -102.4383
6.3143 | -102.4584
6.318 | -102.4794
6.3217 | -102.5024
6.3254 | -102.5229
6.329 | -102.5327
6.3327 | -102.5322
6.3364 | -102.5283
6.3401 | -102.5139
6.3437 | -102.4911
6.3474 | -102.4576
6.3511 | -102.4104
6.3547 | -102.3592
6.3584 | -102.3156
6.3621 | -102.2694
6.3658 | -102.2323
6.3694 | -102.2044
6.3731 | -102.1747
6.3768 | -102.1491
6.3805 | -102.1378
6.3841 | -102.1323
6.3878 | -102.1323
6.3915 | -102.1344
6.3952 | -102.1461
6.3988 | -102.1639
6.4025 | -102.1844
6.4062 | -102.2088
6.4098 | -102.2385
6.4135 | -102.2679
6.4172 | -102.2962
6.4209 | -102.3213
6.4245 | -102.3475
6.4282 | -102.3737
6.4319 | -102.3926
6.4356 | -102.4084
6.4392 | -102.4169
6.4429 | -102.4229
6.4466 | -102.4287
6.4502 | -102.4352
6.4539 | -102.4431
6.4576 | -102.4543
6.4613 | -102.4715
6.4649 | -102.4854
6.4686 | -102.4981
6.4723 | -102.5155
6.476 | -102.5369
6.4796 | -102.5509
6.4833 | -102.5603
6.487 | -102.5712
6.4907 | -102.5811
6.4943 | -102.5924
6.498 | -102.601
6.5017 | -102.6068
6.5053 | -102.6128
6.509 | -102.6082
6.5127 | -102.6028
6.5164 | -102.5974
6.52 | -102.5963
6.5237 | -102.58
6.5274 | -102.5544
6.5311 | -102.5264
6.5347 | -102.4981
6.5384 | -102.4691
6.5421 | -102.4358
6.5458 | -102.4067
6.5494 | -102.3789
6.5531 | -102.3588
6.5568 | -102.3506
6.5604 | -102.3548
6.5641 | -102.3748
6.5678 | -102.4058
6.5715 | -102.4521
6.5751 | -102.505
6.5788 | -102.5721
6.5825 | -102.6433
6.5862 | -102.7099
6.5898 | -102.7754
6.5935 | -102.8333
6.5972 | -102.8821
6.6009 | -102.9212
6.6045 | -102.953
6.6082 | -102.9807
6.6119 | -102.9918
6.6155 | -103.0027
6.6192 | -103.0106
6.6229 | -103.0276
6.6266 | -103.0406
6.6302 | -103.0512
6.6339 | -103.068
6.6376 | -103.0909
6.6413 | -103.1188
6.6449 | -103.1393
6.6486 | -103.1608
6.6523 | -103.1766
6.656 | -103.1879
6.6596 | -103.1933
6.6633 | -103.1848
6.667 | -103.1619
6.6706 | -103.1245
6.6743 | -103.0812
6.678 | -103.0232
6.6817 | -102.9599
6.6853 | -102.8992
6.689 | -102.8358
6.6927 | -102.7718
6.6964 | -102.7187
6.7 | -102.6776
6.7037 | -102.6429
6.7074 | -102.6139
6.7111 | -102.5917
6.7147 | -102.5732
6.7184 | -102.562
6.7221 | -102.5563
6.7257 | -102.5538
6.7294 | -102.5525
6.7331 | -102.5512
6.7368 | -102.5534
6.7404 | -102.5607
6.7441 | -102.5661
6.7478 | -102.5722
6.7515 | -102.582
6.7551 | -102.5954
6.7588 | -102.616
6.7625 | -102.6427
6.7662 | -102.6772
6.7698 | -102.7147
6.7735 | -102.7454
6.7772 | -102.7807
6.7808 | -102.8235
6.7845 | -102.8592
6.7882 | -102.8916
6.7919 | -102.9237
6.7955 | -102.9494
6.7992 | -102.9667
6.8029 | -102.9784
6.8066 | -102.9885
6.8102 | -102.9988
6.8139 | -102.9933
6.8176 | -102.9892
6.8213 | -102.9828
6.8249 | -102.9758
6.8286 | -102.9605
6.8323 | -102.9387
6.8359 | -102.916
6.8396 | -102.8925
6.8433 | -102.8774
6.847 | -102.8573
6.8506 | -102.8383
6.8543 | -102.8285
6.858 | -102.8226
6.8617 | -102.8257
6.8653 | -102.836
6.869 | -102.8501
6.8727 | -102.8687
6.8763 | -102.8785
6.88 | -102.879
6.8837 | -102.8802
6.8874 | -102.8611
6.891 | -102.8555
6.8947 | -102.8441
6.8984 | -102.834
6.9021 | -102.8278
6.9057 | -102.8272
6.9094 | -102.8329
6.9131 | -102.8429
6.9168 | -102.8643
6.9204 | -102.8833
6.9241 | -102.9022
6.9278 | -102.9257
6.9314 | -102.9544
6.9351 | -102.9833
6.9388 | -103.0098
6.9425 | -103.0375
6.9461 | -103.0661
6.9498 | -103.0946
6.9535 | -103.1236
6.9572 | -103.144
6.9608 | -103.163
6.9645 | -103.1832
6.9682 | -103.1979
6.9719 | -103.2087
6.9755 | -103.2167
6.9792 | -103.2256
6.9829 | -103.227
6.9865 | -103.2159
6.9902 | -103.2027
6.9939 | -103.1851
6.9976 | -103.1626
7.0012 | -103.1401
7.0049 | -103.1173
7.0086 | -103.0947
7.0123 | -103.0717
7.0159 | -103.0541
7.0196 | -103.046
7.0233 | -103.0405
7.027 | -103.0408
7.0306 | -103.053
7.0343 | -103.0713
7.038 | -103.0945
7.0416 | -103.1243
7.0453 | -103.1588
7.049 | -103.1937
7.0527 | -103.2238
7.0563 | -103.2557
7.06 | -103.2777
7.0637 | -103.2937
7.0674 | -103.3029
7.071 | -103.303
7.0747 | -103.2988
7.0784 | -103.2916
7.0821 | -103.2872
7.0857 | -103.2846
7.0894 | -103.287
7.0931 | -103.2971
7.0967 | -103.3097
7.1004 | -103.3318
7.1041 | -103.3681
7.1078 | -103.4136
7.1114 | -103.4614
7.1151 | -103.5078
7.1188 | -103.5525
7.1225 | -103.5876
7.1261 | -103.6231
7.1298 | -103.6498
7.1335 | -103.6622
7.1372 | -103.6571
7.1408 | -103.6387
7.1445 | -103.6102
7.1482 | -103.5727
7.1518 | -103.5316
7.1555 | -103.491
7.1592 | -103.4496
7.1629 | -103.4069
7.1665 | -103.3773
7.1702 | -103.3571
7.1739 | -103.3402
7.1776 | -103.3342
7.1812 | -103.3453
7.1849 | -103.366
7.1886 | -103.3915
7.1922 | -103.4206
7.1959 | -103.4516
7.1996 | -103.4813
7.2033 | -103.5097
7.2069 | -103.5258
7.2106 | -103.5317
7.2143 | -103.5332
7.218 | -103.5294
7.2216 | -103.5206
7.2253 | -103.5034
7.229 | -103.4794
7.2327 | -103.4424
7.2363 | -103.4084
7.24 | -103.3807
7.2437 | -103.3544
7.2473 | -103.3305
7.251 | -103.3066
7.2547 | -103.286
7.2584 | -103.2749
7.262 | -103.2719
7.2657 | -103.2832
7.2694 | -103.3038
7.2731 | -103.3326
7.2767 | -103.3748
7.2804 | -103.4306
7.2841 | -103.4889
7.2878 | -103.5499
7.2914 | -103.6159
7.2951 | -103.6847
7.2988 | -103.7545
7.3024 | -103.8213
7.3061 | -103.8817
7.3098 | -103.928
7.3135 | -103.9617
7.3171 | -103.9859
7.3208 | -103.9921
7.3245 | -103.9928
7.3282 | -103.9888
7.3318 | -103.9733
7.3355 | -103.9452
7.3392 | -103.9115
7.3429 | -103.8737
7.3465 | -103.8253
7.3502 | -103.7811
7.3539 | -103.7382
7.3575 | -103.6961
7.3612 | -103.6517
7.3649 | -103.606
7.3686 | -103.557
7.3722 | -103.5057
7.3759 | -103.4497
7.3796 | -103.3934
7.3833 | -103.3344
7.3869 | -103.275
7.3906 | -103.2217
7.3943 | -103.1737
7.398 | -103.1325
7.4016 | -103.1016
7.4053 | -103.0817
7.409 | -103.0707
7.4126 | -103.0641
7.4163 | -103.0947
7.42 | -103.113
7.4237 | -103.1404
7.4273 | -103.1729
7.431 | -103.2177
7.4347 | -103.2643
7.4384 | -103.3093
7.442 | -103.3606
7.4457 | -103.4178
7.4494 | -103.4721
7.4531 | -103.5218
7.4567 | -103.5725
7.4604 | -103.6139
7.4641 | -103.6502
7.4677 | -103.6842
7.4714 | -103.7126
7.4751 | -103.7312
7.4788 | -103.7482
7.4824 | -103.7533
7.4861 | -103.7493
7.4898 | -103.7364
7.4935 | -103.7192
7.4971 | -103.6933
7.5008 | -103.6616
7.5045 | -103.6251
7.5081 | -103.5829
7.5118 | -103.5395
7.5155 | -103.4933
7.5192 | -103.4524
7.5228 | -103.4181
7.5265 | -103.3828
7.5302 | -103.349
7.5339 | -103.319
7.5375 | -103.2969
7.5412 | -103.2852
7.5449 | -103.2821
7.5486 | -103.2846
7.5522 | -103.2975
7.5559 | -103.322
7.5596 | -103.3524
7.5632 | -103.389
7.5669 | -103.432
7.5706 | -103.4763
7.5743 | -103.5186
7.5779 | -103.5584
7.5816 | -103.5876
7.5853 | -103.6097
7.589 | -103.6247
7.5926 | -103.6297
7.5963 | -103.6354
7.6 | -103.6391
7.6037 | -103.6288
7.6073 | -103.6181
7.611 | -103.6114
7.6147 | -103.6
7.6183 | -103.59
7.622 | -103.5829
7.6257 | -103.5736
7.6294 | -103.5759
7.633 | -103.5822
7.6367 | -103.592
7.6404 | -103.608
7.6441 | -103.6272
7.6477 | -103.6469
7.6514 | -103.6662
7.6551 | -103.682
7.6588 | -103.6987
7.6624 | -103.7142
7.6661 | -103.7231
7.6698 | -103.7204
7.6734 | -103.7155
7.6771 | -103.7068
7.6808 | -103.6917
7.6845 | -103.675
7.6881 | -103.6564
7.6918 | -103.6407
7.6955 | -103.6297
7.6992 | -103.6198
7.7028 | -103.6105
7.7065 | -103.603
7.7102 | -103.6008
7.7139 | -103.6058
7.7175 | -103.6221
7.7212 | -103.643
7.7249 | -103.6726
7.7285 | -103.7104
7.7322 | -103.7538
7.7359 | -103.8015
7.7396 | -103.8578
7.7432 | -103.9094
7.7469 | -103.964
7.7506 | -104.0219
7.7543 | -104.0751
7.7579 | -104.1212
7.7616 | -104.1624
7.7653 | -104.2003
7.769 | -104.2265
7.7726 | -104.2414
7.7763 | -104.2571
7.78 | -104.2664
7.7836 | -104.2672
7.7873 | -104.2663
7.791 | -104.2703
7.7947 | -104.2705
7.7983 | -104.2692
7.802 | -104.2687
7.8057 | -104.2577
7.8094 | -104.2447
7.813 | -104.2296
7.8167 | -104.2135
7.8204 | -104.1949
7.8241 | -104.167
7.8277 | -104.1349
7.8314 | -104.1042
7.8351 | -104.0727
7.8387 | -104.0426
7.8424 | -104.0166
7.8461 | -103.9954
7.8498 | -103.9772
7.8534 | -103.9595
7.8571 | -103.9471
7.8608 | -103.9356
7.8645 | -103.9325
7.8681 | -103.9339
7.8718 | -103.9412
7.8755 | -103.9474
7.8792 | -103.9534
7.8828 | -103.9694
7.8865 | -103.986
7.8902 | -103.9964
7.8938 | -104.0027
7.8975 | -104.0109
7.9012 | -104.0115
7.9049 | -104.0059
7.9085 | -104.0034
7.9122 | -103.9999
7.9159 | -103.997
7.9196 | -103.9948
7.9232 | -103.9925
7.9269 | -103.9914
7.9306 | -103.9916
7.9342 | -103.998
7.9379 | -104.0149
7.9416 | -104.0317
7.9453 | -104.0484
7.9489 | -104.0704
7.9526 | -104.0923
7.9563 | -104.1126
7.96 | -104.1324
7.9636 | -104.1519
7.9673 | -104.173
7.971 | -104.1891
7.9747 | -104.2013
7.9783 | -104.214
7.982 | -104.2232
7.9857 | -104.2343
7.9893 | -104.2439
7.993 | -104.2495
7.9967 | -104.2551
8.0004 | -104.2522
8.004 | -104.2469
8.0077 | -104.244
8.0114 | -104.2387
8.0151 | -104.2284
8.0187 | -104.2156
8.0224 | -104.1977
8.0261 | -104.1718
8.0298 | -104.1423
8.0334 | -104.1119
8.0371 | -104.0845
8.0408 | -104.0585
8.0444 | -104.036
8.0481 | -104.0203
8.0518 | -104.0115
8.0555 | -104.0084
8.0591 | -104.019
8.0628 | -104.04
8.0665 | -104.0696
8.0702 | -104.1034
8.0738 | -104.1409
8.0775 | -104.1893
8.0812 | -104.2396
8.0849 | -104.2913
8.0885 | -104.3445
8.0922 | -104.3953
8.0959 | -104.444
8.0995 | -104.4894
8.1032 | -104.5238
8.1069 | -104.5472
8.1106 | -104.5663
8.1142 | -104.5799
8.1179 | -104.5869
8.1216 | -104.5898
8.1253 | -104.5866
8.1289 | -104.5722
8.1326 | -104.5575
8.1363 | -104.5395
8.14 | -104.5144
8.1436 | -104.4896
8.1473 | -104.4734
8.151 | -104.4558
8.1546 | -104.4336
8.1583 | -104.4125
8.162 | -104.3887
8.1657 | -104.3628
8.1693 | -104.3378
8.173 | -104.312
8.1767 | -104.287
8.1804 | -104.2595
8.184 | -104.2303
8.1877 | -104.2063
8.1914 | -104.1816
8.1951 | -104.1617
8.1987 | -104.1442
8.2024 | -104.1332
8.2061 | -104.126
8.2097 | -104.1241
8.2134 | -104.1304
8.2171 | -104.1393
8.2208 | -104.1526
8.2244 | -104.1733
8.2281 | -104.1978
8.2318 | -104.2207
8.2355 | -104.2452
8.2391 | -104.2616
8.2428 | -104.2779
8.2465 | -104.2934
8.2501 | -104.3117
8.2538 | -104.3202
8.2575 | -104.3228
8.2612 | -104.3234
8.2648 | -104.3208
8.2685 | -104.3092
8.2722 | -104.3014
8.2759 | -104.2965
8.2795 | -104.2858
8.2832 | -104.2756
8.2869 | -104.2699
8.2906 | -104.2642
8.2942 | -104.2592
8.2979 | -104.2561
8.3016 | -104.2572
8.3052 | -104.2605
8.3089 | -104.2625
8.3126 | -104.2716
8.3163 | -104.2839
8.3199 | -104.297
8.3236 | -104.3127
8.3273 | -104.328
8.331 | -104.3427
8.3346 | -104.3612
8.3383 | -104.3773
8.342 | -104.3855
8.3457 | -104.3909
8.3493 | -104.3907
8.353 | -104.3831
8.3567 | -104.3665
8.3603 | -104.3494
8.364 | -104.3296
8.3677 | -104.307
8.3714 | -104.28
8.375 | -104.2581
8.3787 | -104.2354
8.3824 | -104.2157
8.3861 | -104.1974
8.3897 | -104.183
8.3934 | -104.1721
8.3971 | -104.1676
8.4008 | -104.167
8.4044 | -104.1667
8.4081 | -104.1686
8.4118 | -104.174
8.4154 | -104.1828
8.4191 | -104.1914
8.4228 | -104.2012
8.4265 | -104.2079
8.4301 | -104.2135
8.4338 | -104.2184
8.4375 | -104.2221
8.4412 | -104.2245
8.4448 | -104.2304
8.4485 | -104.2378
8.4522 | -104.2443
8.4559 | -104.2463
8.4595 | -104.2525
8.4632 | -104.2624
8.4669 | -104.2731
8.4705 | -104.2866
8.4742 | -104.294
8.4779 | -104.2879
8.4816 | -104.2893
8.4852 | -104.2909
8.4889 | -104.2858
8.4926 | -104.2736
8.4963 | -104.2587
8.4999 | -104.2426
8.5036 | -104.2276
8.5073 | -104.2114
8.511 | -104.1949
8.5146 | -104.1795
8.5183 | -104.1641
8.522 | -104.1458
8.5256 | -104.1315
8.5293 | -104.1224
8.533 | -104.1099
8.5367 | -104.0958
8.5403 | -104.0808
8.544 | -104.0681
8.5477 | -104.053
8.5514 | -104.0435
8.555 | -104.0289
8.5587 | -104.0075
8.5624 | -103.9906
8.5661 | -103.9762
8.5697 | -103.9623
8.5734 | -103.9496
8.5771 | -103.9448
8.5807 | -103.9437
8.5844 | -103.9471
8.5881 | -103.9514
8.5918 | -103.959
8.5954 | -103.9708
8.5991 | -103.9921
8.6028 | -104.0159
8.6065 | -104.0423
8.6101 | -104.0717
8.6138 | -104.1019
8.6175 | -104.1394
8.6211 | -104.179
8.6248 | -104.2178
8.6285 | -104.255
8.6322 | -104.2884
8.6358 | -104.3198
8.6395 | -104.3465
8.6432 | -104.3694
8.6469 | -104.3908
8.6505 | -104.4134
8.6542 | -104.4292
8.6579 | -104.4395
8.6616 | -104.4437
8.6652 | -104.4414
8.6689 | -104.4382
8.6726 | -104.4312
8.6762 | -104.4224
8.6799 | -104.4095
8.6836 | -104.3984
8.6873 | -104.3922
8.6909 | -104.3827
8.6946 | -104.3721
8.6983 | -104.3627
8.702 | -104.3528
8.7056 | -104.3364
8.7093 | -104.3201
8.713 | -104.3014
8.7167 | -104.288
8.7203 | -104.277
8.724 | -104.2651
8.7277 | -104.2528
8.7313 | -104.24
8.735 | -104.2297
8.7387 | -104.2206
8.7424 | -104.2128
8.746 | -104.2091
8.7497 | -104.2134
8.7534 | -104.2184
8.7571 | -104.2234
8.7607 | -104.231
8.7644 | -104.2388
8.7681 | -104.2477
8.7718 | -104.2563
8.7754 | -104.2688
8.7791 | -104.2805
8.7828 | -104.2924
8.7864 | -104.3067
8.7901 | -104.3266
8.7938 | -104.348
8.7975 | -104.3671
8.8011 | -104.3881
8.8048 | -104.4107
8.8085 | -104.4338
8.8122 | -104.4572
8.8158 | -104.473
8.8195 | -104.4886
8.8232 | -104.5057
8.8269 | -104.5223
8.8305 | -104.5405
8.8342 | -104.5606
8.8379 | -104.5781
8.8415 | -104.5937
8.8452 | -104.6043
8.8489 | -104.6149
8.8526 | -104.6243
8.8562 | -104.6346
8.8599 | -104.6468
8.8636 | -104.6575
8.8673 | -104.6645
8.8709 | -104.6719
8.8746 | -104.6771
8.8783 | -104.6793
8.882 | -104.6831
8.8856 | -104.6897
8.8893 | -104.6953
8.893 | -104.6989
8.8966 | -104.7061
8.9003 | -104.7083
8.904 | -104.7048
8.9077 | -104.7026
8.9113 | -104.7022
8.915 | -104.6961
8.9187 | -104.6841
8.9224 | -104.6747
8.926 | -104.6655
8.9297 | -104.6483
8.9334 | -104.6244
8.937 | -104.6052
8.9407 | -104.593
8.9444 | -104.5865
8.9481 | -104.5722
8.9517 | -104.5625
8.9554 | -104.5552
8.9591 | -104.5436
8.9628 | -104.5296
8.9664 | -104.5208
8.9701 | -104.5098
8.9738 | -104.498
8.9775 | -104.4769
8.9811 | -104.4545
8.9848 | -104.4312
8.9885 | -104.4055
8.9921 | -104.3761
8.9958 | -104.3462
8.9995 | -104.3178
9.0032 | -104.2897
9.0068 | -104.267
9.0105 | -104.2459
9.0142 | -104.2296
9.0179 | -104.2182
9.0215 | -104.2117
9.0252 | -104.2072
9.0289 | -104.2048
9.0326 | -104.2083
9.0362 | -104.2188
9.0399 | -104.2316
9.0436 | -104.2417
9.0472 | -104.2534
9.0509 | -104.2704
9.0546 | -104.2864
9.0583 | -104.2992
9.0619 | -104.3138
9.0656 | -104.3322
9.0693 | -104.3395
9.073 | -104.3394
9.0766 | -104.3397
9.0803 | -104.3448
9.084 | -104.3441
9.0877 | -104.3423
9.0913 | -104.3448
9.095 | -104.3489
9.0987 | -104.3488
9.1023 | -104.3534
9.106 | -104.3637
9.1097 | -104.3811
9.1134 | -104.4006
9.117 | -104.4205
9.1207 | -104.4459
9.1244 | -104.4801
9.1281 | -104.5137
9.1317 | -104.5473
9.1354 | -104.5851
9.1391 | -104.6226
9.1428 | -104.6525
9.1464 | -104.6809
9.1501 | -104.7037
9.1538 | -104.7176
9.1574 | -104.7319
9.1611 | -104.7507
9.1648 | -104.7648
9.1685 | -104.7797
9.1721 | -104.7859
9.1758 | -104.7885
9.1795 | -104.795
9.1832 | -104.8025
9.1868 | -104.8037
9.1905 | -104.8014
9.1942 | -104.8022
9.1979 | -104.8054
9.2015 | -104.8111
9.2052 | -104.8229
9.2089 | -104.8347
9.2125 | -104.8459
9.2162 | -104.8581
9.2199 | -104.8731
9.2236 | -104.892
9.2272 | -104.9142
9.2309 | -104.9309
9.2346 | -104.9483
9.2383 | -104.9658
9.2419 | -104.981
9.2456 | -104.9939
9.2493 | -105.0028
9.253 | -105.0059
9.2566 | -105.0086
9.2603 | -105.0059
9.264 | -105.0025
9.2676 | -104.9928
9.2713 | -104.973
9.275 | -104.9519
9.2787 | -104.926
9.2823 | -104.8966
9.286 | -104.8664
9.2897 | -104.8361
9.2934 | -104.7991
9.297 | -104.7584
9.3007 | -104.7204
9.3044 | -104.6819
9.308 | -104.6483
9.3117 | -104.6231
9.3154 | -104.6004
9.3191 | -104.5717
9.3227 | -104.5476
9.3264 | -104.5276
9.3301 | -104.5109
9.3338 | -104.4969
9.3374 | -104.4823
9.3411 | -104.4684
9.3448 | -104.4598
9.3485 | -104.4529
9.3521 | -104.4496
9.3558 | -104.4447
9.3595 | -104.4387
9.3631 | -104.4334
9.3668 | -104.4317
9.3705 | -104.4307
9.3742 | -104.4224
9.3778 | -104.4174
9.3815 | -104.4132
9.3852 | -104.4082
9.3889 | -104.4082
9.3925 | -104.4053
9.3962 | -104.4027
9.3999 | -104.403
9.4036 | -104.4108
9.4072 | -104.4251
9.4109 | -104.4395
9.4146 | -104.4559
9.4182 | -104.4738
9.4219 | -104.4919
9.4256 | -104.5064
9.4293 | -104.5245
9.4329 | -104.5446
9.4366 | -104.5648
9.4403 | -104.5854
9.444 | -104.6048
9.4476 | -104.6237
9.4513 | -104.6416
9.455 | -104.6671
9.4587 | -104.6881
9.4623 | -104.7094
9.466 | -104.7307
9.4697 | -104.7483
9.4733 | -104.7673
9.477 | -104.7871
9.4807 | -104.8084
9.4844 | -104.8262
9.488 | -104.8408
9.4917 | -104.8589
9.4954 | -104.8741
9.4991 | -104.8861
9.5027 | -104.8992
9.5064 | -104.9073
9.5101 | -104.9213
9.5138 | -104.9288
9.5174 | -104.9304
9.5211 | -104.9345
9.5248 | -104.9383
9.5284 | -104.9407
9.5321 | -104.9427
9.5358 | -104.9399
9.5395 | -104.9401
9.5431 | -104.9359
9.5468 | -104.9334
9.5505 | -104.9312
9.5542 | -104.9267
9.5578 | -104.9201
9.5615 | -104.9117
9.5652 | -104.9046
9.5689 | -104.8945
9.5725 | -104.8835
9.5762 | -104.8703
9.5799 | -104.857
9.5835 | -104.8436
9.5872 | -104.8292
9.5909 | -104.815
9.5946 | -104.798
9.5982 | -104.7845
9.6019 | -104.7684
9.6056 | -104.7503
9.6093 | -104.7362
9.6129 | -104.7235
9.6166 | -104.7166
9.6203 | -104.7101
9.624 | -104.7018
9.6276 | -104.6985
9.6313 | -104.698
9.635 | -104.6966
9.6386 | -104.7003
9.6423 | -104.7059
9.646 | -104.7156
9.6497 | -104.7222
9.6533 | -104.7308
9.657 | -104.7426
9.6607 | -104.749
9.6644 | -104.7507
9.668 | -104.7523
9.6717 | -104.7546
9.6754 | -104.7591
9.679 | -104.7645
9.6827 | -104.7647
9.6864 | -104.764
9.6901 | -104.7643
9.6937 | -104.7687
9.6974 | -104.7769
9.7011 | -104.7882
9.7048 | -104.8036
9.7084 | -104.8208
9.7121 | -104.8358
9.7158 | -104.8486
9.7195 | -104.8656
9.7231 | -104.8857
9.7268 | -104.905
9.7305 | -104.9229
9.7341 | -104.9353
9.7378 | -104.9483
9.7415 | -104.9568
9.7452 | -104.9638
9.7488 | -104.9687
9.7525 | -104.9663
9.7562 | -104.9587
9.7599 | -104.9487
9.7635 | -104.9343
9.7672 | -104.9229
9.7709 | -104.9048
9.7746 | -104.8849
9.7782 | -104.87
9.7819 | -104.8491
9.7856 | -104.826
9.7892 | -104.8058
9.7929 | -104.7913
9.7966 | -104.781
9.8003 | -104.7745
9.8039 | -104.7692
9.8076 | -104.7674
9.8113 | -104.7665
9.815 | -104.7717
9.8186 | -104.7794
9.8223 | -104.787
9.826 | -104.7995
9.8297 | -104.8149
9.8333 | -104.827
9.837 | -104.8397
9.8407 | -104.8495
9.8443 | -104.8574
9.848 | -104.8609
9.8517 | -104.8617
9.8554 | -104.8646
9.859 | -104.8709
9.8627 | -104.8747
9.8664 | -104.8779
9.8701 | -104.8814
9.8737 | -104.8864
9.8774 | -104.9002
9.8811 | -104.9123
9.8848 | -104.9261
9.8884 | -104.9415
9.8921 | -104.9579
9.8958 | -104.9801
9.8994 | -105.0051
9.9031 | -105.0271
9.9068 | -105.0508
9.9105 | -105.0748
9.9141 | -105.0989
9.9178 | -105.1194
9.9215 | -105.1368
9.9252 | -105.156
9.9288 | -105.1737
9.9325 | -105.1896
9.9362 | -105.2033
9.9398 | -105.2086
9.9435 | -105.2105
9.9472 | -105.2156
9.9509 | -105.2169
9.9545 | -105.2155
9.9582 | -105.2051
9.9619 | -105.1921
9.9656 | -105.1811
9.9692 | -105.1661
9.9729 | -105.1467
9.9766 | -105.1253
9.9803 | -105.1022
9.9839 | -105.0841
9.9876 | -105.0632
9.9913 | -105.0448
9.9949 | -105.0253
9.9986 | -104.9999
10.0023 | -104.9756
10.006 | -104.9545
10.0096 | -104.934
10.0133 | -104.9207
10.017 | -104.9106
10.0207 | -104.9
10.0243 | -104.8947
10.028 | -104.8929
10.0317 | -104.8892
10.0354 | -104.8828
10.039 | -104.8748
10.0427 | -104.8703
10.0464 | -104.8676
10.05 | -104.8644
10.0537 | -104.8612
10.0574 | -104.8587
10.0611 | -104.8589
10.0647 | -104.8571
10.0684 | -104.8631
10.0721 | -104.862
10.0758 | -104.8636
10.0794 | -104.8667
10.0831 | -104.8708
10.0868 | -104.8802
10.0905 | -104.893
10.0941 | -104.9039
10.0978 | -104.9144
10.1015 | -104.9281
10.1051 | -104.9438
10.1088 | -104.9572
10.1125 | -104.9728
10.1162 | -104.9907
10.1198 | -105.0078
10.1235 | -105.0249
10.1272 | -105.0389
10.1309 | -105.0487
10.1345 | -105.0596
10.1382 | -105.0693
10.1419 | -105.083
10.1456 | -105.0951
10.1492 | -105.1008
10.1529 | -105.1107
10.1566 | -105.1208
10.1602 | -105.1255
10.1639 | -105.1307
10.1676 | -105.1392
10.1713 | -105.1463
10.1749 | -105.1564
10.1786 | -105.169
10.1823 | -105.1802
10.186 | -105.1913
10.1896 | -105.1969
10.1933 | -105.2031
10.197 | -105.2103
10.2007 | -105.2138
10.2043 | -105.2154
10.208 | -105.2148
10.2117 | -105.2172
10.2153 | -105.2195
10.219 | -105.224
10.2227 | -105.2308
10.2264 | -105.2363
10.23 | -105.2392
10.2337 | -105.2437
10.2374 | -105.2489
10.2411 | -105.2538
10.2447 | -105.2601
10.2484 | -105.2728
10.2521 | -105.2846
10.2558 | -105.293
10.2594 | -105.3062
10.2631 | -105.3205
10.2668 | -105.3345
10.2704 | -105.3507
10.2741 | -105.3679
10.2778 | -105.3886
10.2815 | -105.4085
10.2851 | -105.4253
10.2888 | -105.4393
10.2925 | -105.4495
10.2962 | -105.459
10.2998 | -105.4649
10.3035 | -105.4604
10.3072 | -105.4549
10.3109 | -105.4493
10.3145 | -105.4407
10.3182 | -105.4278
10.3219 | -105.4089
10.3255 | -105.3821
10.3292 | -105.3573
10.3329 | -105.333
10.3366 | -105.3
10.3402 | -105.2663
10.3439 | -105.2289
10.3476 | -105.1904
10.3513 | -105.1551
10.3549 | -105.1194
10.3586 | -105.0849
10.3623 | -105.0541
10.3659 | -105.0232
10.3696 | -104.9937
10.3733 | -104.9717
10.377 | -104.9532
10.3806 | -104.9379
10.3843 | -104.926
10.388 | -104.9228
10.3917 | -104.9225
10.3953 | -104.9248
10.399 | -104.9264
10.4027 | -104.9351
10.4064 | -104.9457
10.41 | -104.9524
10.4137 | -104.9615
10.4174 | -104.9709
10.421 | -104.9773
10.4247 | -104.9813
10.4284 | -104.983
10.4321 | -104.9872
10.4357 | -104.9933
10.4394 | -104.9978
10.4431 | -104.997
10.4468 | -105.0003
10.4504 | -105.0035
10.4541 | -105.0061
10.4578 | -105.0119
10.4615 | -105.0173
10.4651 | -105.0169
10.4688 | -105.0169
10.4725 | -105.0211
10.4761 | -105.0273
10.4798 | -105.0294
10.4835 | -105.0319
10.4872 | -105.0397
10.4908 | -105.0455
10.4945 | -105.0525
10.4982 | -105.0592
10.5019 | -105.0653
10.5055 | -105.0741
10.5092 | -105.0864
10.5129 | -105.1003
10.5166 | -105.1143
10.5202 | -105.1348
10.5239 | -105.1543
10.5276 | -105.172
10.5312 | -105.1929
10.5349 | -105.2152
10.5386 | -105.2328
10.5423 | -105.2498
10.5459 | -105.2673
10.5496 | -105.2857
10.5533 | -105.3023
10.557 | -105.3204
10.5606 | -105.3392
10.5643 | -105.3527
10.568 | -105.3666
10.5717 | -105.3806
10.5753 | -105.3924
10.579 | -105.4004
10.5827 | -105.4008
10.5863 | -105.398
10.59 | -105.3945
10.5937 | -105.3863
10.5974 | -105.3658
10.601 | -105.3515
10.6047 | -105.3364
10.6084 | -105.3209
10.6121 | -105.2999
10.6157 | -105.2801
10.6194 | -105.265
10.6231 | -105.248
10.6268 | -105.2309
10.6304 | -105.2146
10.6341 | -105.2024
10.6378 | -105.1925
10.6414 | -105.1836
10.6451 | -105.1801
10.6488 | -105.1794
10.6525 | -105.179
10.6561 | -105.1821
10.6598 | -105.188
10.6635 | -105.1931
10.6672 | -105.1976
10.6708 | -105.2096
10.6745 | -105.2195
10.6782 | -105.229
10.6818 | -105.2446
10.6855 | -105.2565
10.6892 | -105.2651
10.6929 | -105.2733
10.6965 | -105.2838
10.7002 | -105.2928
10.7039 | -105.2994
10.7076 | -105.3065
10.7112 | -105.3123
10.7149 | -105.3162
10.7186 | -105.3199
10.7223 | -105.3225
10.7259 | -105.3282
10.7296 | -105.3362
10.7333 | -105.3434
10.7369 | -105.3514
10.7406 | -105.3599
10.7443 | -105.3698
10.748 | -105.3846
10.7516 | -105.3959
10.7553 | -105.4037
10.759 | -105.4127
10.7627 | -105.4214
10.7663 | -105.4286
10.77 | -105.4324
10.7737 | -105.434
10.7774 | -105.4357
10.781 | -105.4359
10.7847 | -105.4342
10.7884 | -105.4286
10.792 | -105.4191
10.7957 | -105.4095
10.7994 | -105.4018
10.8031 | -105.3936
10.8067 | -105.3852
10.8104 | -105.3787
10.8141 | -105.3701
10.8178 | -105.3603
10.8214 | -105.3529
10.8251 | -105.3445
10.8288 | -105.3315
10.8325 | -105.3221
10.8361 | -105.3171
10.8398 | -105.3108
10.8435 | -105.3059
10.8471 | -105.3061
10.8508 | -105.305
10.8545 | -105.3044
10.8582 | -105.3039
10.8618 | -105.304
10.8655 | -105.3041
10.8692 | -105.3093
10.8729 | -105.3141
10.8765 | -105.3173
10.8802 | -105.3185
10.8839 | -105.3197
10.8876 | -105.3194
10.8912 | -105.3177
10.8949 | -105.3174
10.8986 | -105.3131
10.9022 | -105.3032
10.9059 | -105.2911
10.9096 | -105.2795
10.9133 | -105.2691
10.9169 | -105.2589
10.9206 | -105.2453
10.9243 | -105.2315
10.928 | -105.219
10.9316 | -105.2038
10.9353 | -105.189
10.939 | -105.1771
10.9427 | -105.168
10.9463 | -105.1623
10.95 | -105.1562
10.9537 | -105.1485
10.9573 | -105.1432
10.961 | -105.1403
10.9647 | -105.142
10.9684 | -105.1499
10.972 | -105.1624
10.9757 | -105.178
10.9794 | -105.1958
10.9831 | -105.2183
10.9867 | -105.242
10.9904 | -105.2653
10.9941 | -105.2905
10.9977 | -105.318
11.0014 | -105.3486
11.0051 | -105.379
11.0088 | -105.4077
11.0124 | -105.4339
11.0161 | -105.4632
11.0198 | -105.4928
11.0235 | -105.5197
11.0271 | -105.5412
11.0308 | -105.5581
11.0345 | -105.5787
11.0382 | -105.6013
11.0418 | -105.616
11.0455 | -105.6251
11.0492 | -105.6352
11.0528 | -105.6447
11.0565 | -105.6554
11.0602 | -105.6654
11.0639 | -105.6721
11.0675 | -105.6761
11.0712 | -105.6822
11.0749 | -105.6853
11.0786 | -105.6888
11.0822 | -105.6921
11.0859 | -105.696
11.0896 | -105.7002
11.0933 | -105.7029
11.0969 | -105.7055
11.1006 | -105.7081
11.1043 | -105.711
11.1079 | -105.7164
11.1116 | -105.7237
11.1153 | -105.7291
11.119 | -105.7311
11.1226 | -105.7356
11.1263 | -105.7408
11.13 | -105.7452
11.1337 | -105.749
11.1373 | -105.7512
11.141 | -105.7541
11.1447 | -105.7539
11.1484 | -105.7609
11.152 | -105.7651
11.1557 | -105.7669
11.1594 | -105.7685
11.163 | -105.7682
11.1667 | -105.7669
11.1704 | -105.7704
11.1741 | -105.7729
11.1777 | -105.7697
11.1814 | -105.7661
11.1851 | -105.7629
11.1888 | -105.7586
11.1924 | -105.7574
11.1961 | -105.75
11.1998 | -105.7386
11.2035 | -105.7287
11.2071 | -105.723
11.2108 | -105.7149
11.2145 | -105.702
11.2181 | -105.6892
11.2218 | -105.6799
11.2255 | -105.67
11.2292 | -105.6579
11.2328 | -105.6414
11.2365 | -105.6279
11.2402 | -105.6176
11.2439 | -105.6054
11.2475 | -105.5937
11.2512 | -105.5828
11.2549 | -105.5708
11.2586 | -105.558
11.2622 | -105.5458
11.2659 | -105.5323
11.2696 | -105.5208
11.2732 | -105.5077
11.2769 | -105.4928
11.2806 | -105.4824
11.2843 | -105.4697
11.2879 | -105.4532
11.2916 | -105.44
11.2953 | -105.426
11.299 | -105.4141
11.3026 | -105.4051
11.3063 | -105.3977
11.31 | -105.3937
11.3137 | -105.3873
11.3173 | -105.3845
11.321 | -105.3849
11.3247 | -105.3894
11.3283 | -105.3967
11.332 | -105.406
11.3357 | -105.4198
11.3394 | -105.4352
11.343 | -105.4509
11.3467 | -105.4669
11.3504 | -105.49
11.3541 | -105.5142
11.3577 | -105.5404
11.3614 | -105.5687
11.3651 | -105.5984
11.3688 | -105.6259
11.3724 | -105.6535
11.3761 | -105.6796
11.3798 | -105.7059
11.3834 | -105.7324
11.3871 | -105.756
11.3908 | -105.7812
11.3945 | -105.8066
11.3981 | -105.8298
11.4018 | -105.845
11.4055 | -105.8532
11.4092 | -105.8634
11.4128 | -105.871
11.4165 | -105.8764
11.4202 | -105.8804
11.4238 | -105.8808
11.4275 | -105.879
11.4312 | -105.8781
11.4349 | -105.8744
11.4385 | -105.8675
11.4422 | -105.8617
11.4459 | -105.8543
11.4496 | -105.8461
11.4532 | -105.8413
11.4569 | -105.8367
11.4606 | -105.8347
11.4643 | -105.8315
11.4679 | -105.8282
11.4716 | -105.8273
11.4753 | -105.8303
11.4789 | -105.8326
11.4826 | -105.8362
11.4863 | -105.842
11.49 | -105.8504
11.4936 | -105.857
11.4973 | -105.8645
11.501 | -105.8714
11.5047 | -105.881
11.5083 | -105.8901
11.512 | -105.8964
11.5157 | -105.8932
11.5194 | -105.8916
11.523 | -105.8853
11.5267 | -105.8754
11.5304 | -105.8651
11.534 | -105.8569
11.5377 | -105.8432
11.5414 | -105.8269
11.5451 | -105.8103
11.5487 | -105.7883
11.5524 | -105.769
11.5561 | -105.7515
11.5598 | -105.7323
11.5634 | -105.7149
11.5671 | -105.7065
11.5708 | -105.7003
11.5745 | -105.6943
11.5781 | -105.6935
11.5818 | -105.6927
11.5855 | -105.6922
11.5891 | -105.699
11.5928 | -105.7074
11.5965 | -105.7152
11.6002 | -105.7247
11.6038 | -105.7395
11.6075 | -105.7568
11.6112 | -105.7705
11.6149 | -105.7827
11.6185 | -105.7945
11.6222 | -105.8049
11.6259 | -105.811
11.6296 | -105.8139
11.6332 | -105.8167
11.6369 | -105.8178
11.6406 | -105.8182
11.6442 | -105.8143
11.6479 | -105.8114
11.6516 | -105.8051
11.6553 | -105.7885
11.6589 | -105.7767
11.6626 | -105.768
11.6663 | -105.755
11.67 | -105.7385
11.6736 | -105.7245
11.6773 | -105.7128
11.681 | -105.6973
11.6846 | -105.6815
11.6883 | -105.6628
11.692 | -105.6442
11.6957 | -105.6275
11.6993 | -105.612
11.703 | -105.5952
11.7067 | -105.5816
11.7104 | -105.5675
11.714 | -105.5515
11.7177 | -105.5413
11.7214 | -105.5308
11.7251 | -105.5201
11.7287 | -105.5132
11.7324 | -105.5066
11.7361 | -105.4996
11.7397 | -105.4976
11.7434 | -105.4949
11.7471 | -105.4876
11.7508 | -105.4861
11.7544 | -105.4917
11.7581 | -105.4909
11.7618 | -105.4894
11.7655 | -105.4919
11.7691 | -105.4973
11.7728 | -105.5031
11.7765 | -105.507
11.7802 | -105.5154
11.7838 | -105.5266
11.7875 | -105.5355
11.7912 | -105.5443
11.7948 | -105.5571
11.7985 | -105.5713
11.8022 | -105.5832
11.8059 | -105.5941
11.8095 | -105.6011
11.8132 | -105.6099
11.8169 | -105.6156
11.8206 | -105.6183
11.8242 | -105.6225
11.8279 | -105.6291
11.8316 | -105.6324
11.8353 | -105.6344
11.8389 | -105.6364
11.8426 | -105.6349
11.8463 | -105.6302
11.8499 | -105.6229
11.8536 | -105.6173
11.8573 | -105.6131
11.861 | -105.6069
11.8646 | -105.6018
11.8683 | -105.5976
11.872 | -105.5899
11.8757 | -105.5859
11.8793 | -105.5838
11.883 | -105.5856
11.8867 | -105.5908
11.8904 | -105.5939
11.894 | -105.5963
11.8977 | -105.6003
11.9014 | -105.6044
11.905 | -105.6079
11.9087 | -105.6121
11.9124 | -105.619
11.9161 | -105.6278
11.9197 | -105.6426
11.9234 | -105.6531
11.9271 | -105.6599
11.9308 | -105.6749
11.9344 | -105.6904
11.9381 | -105.7028
11.9418 | -105.7118
11.9455 | -105.7192
11.9491 | -105.7261
11.9528 | -105.7317
11.9565 | -105.7353
11.9601 | -105.7398
11.9638 | -105.7433
11.9675 | -105.7427
11.9712 | -105.7408
11.9748 | -105.7381
11.9785 | -105.7311
11.9822 | -105.7244
11.9859 | -105.7226
11.9895 | -105.7194
11.9932 | -105.7142
11.9969 | -105.705
12.0006 | -105.6929
12.0042 | -105.6837
12.0079 | -105.6744
12.0116 | -105.6678
12.0152 | -105.6631
12.0189 | -105.6562
12.0226 | -105.65
12.0263 | -105.6469
12.0299 | -105.6457
12.0336 | -105.6475
12.0373 | -105.6494
12.041 | -105.6518
12.0446 | -105.6601
12.0483 | -105.6698
12.052 | -105.6786
12.0556 | -105.6869
12.0593 | -105.6993
12.063 | -105.7154
12.0667 | -105.7375
12.0703 | -105.76
12.074 | -105.7864
12.0777 | -105.8119
12.0814 | -105.8346
12.085 | -105.855
12.0887 | -105.8777
12.0924 | -105.9002
12.0961 | -105.9273
12.0997 | -105.9539
12.1034 | -105.9759
12.1071 | -105.9957
12.1107 | -106.0115
12.1144 | -106.0272
12.1181 | -106.0447
12.1218 | -106.0603
12.1254 | -106.0749
12.1291 | -106.0887
12.1328 | -106.0934
12.1365 | -106.0931
12.1401 | -106.0924
12.1438 | -106.0914
12.1475 | -106.0868
12.1512 | -106.0831
12.1548 | -106.0742
12.1585 | -106.0589
12.1622 | -106.0447
12.1658 | -106.0298
12.1695 | -106.0132
12.1732 | -105.9948
12.1769 | -105.9704
12.1805 | -105.9424
12.1842 | -105.9126
12.1879 | -105.8738
12.1916 | -105.8485
12.1952 | -105.8263
12.1989 | -105.8029
12.2026 | -105.7749
12.2063 | -105.7482
12.2099 | -105.7244
12.2136 | -105.7023
12.2173 | -105.6812
12.2209 | -105.6606
12.2246 | -105.6428
12.2283 | -105.6293
12.232 | -105.624
12.2356 | -105.6191
12.2393 | -105.6158
12.243 | -105.6144
12.2467 | -105.616
12.2503 | -105.6228
12.254 | -105.6372
12.2577 | -105.6533
12.2614 | -105.673
12.265 | -105.6936
12.2687 | -105.7151
12.2724 | -105.7362
12.276 | -105.7558
12.2797 | -105.7768
12.2834 | -105.8013
12.2871 | -105.8305
12.2907 | -105.8575
12.2944 | -105.8777
12.2981 | -105.8972
12.3018 | -105.9164
12.3054 | -105.9291
12.3091 | -105.9436
12.3128 | -105.9544
12.3165 | -105.9589
12.3201 | -105.9615
12.3238 | -105.9607
12.3275 | -105.9561
12.3311 | -105.9475
12.3348 | -105.9342
12.3385 | -105.9186
12.3422 | -105.9035
12.3458 | -105.8874
12.3495 | -105.867
12.3532 | -105.8472
12.3569 | -105.8286
12.3605 | -105.806
12.3642 | -105.7855
12.3679 | -105.7628
12.3716 | -105.7399
12.3752 | -105.7178
12.3789 | -105.6995
12.3826 | -105.6862
12.3862 | -105.6729
12.3899 | -105.6609
12.3936 | -105.6526
12.3973 | -105.6476
12.4009 | -105.6459
12.4046 | -105.6436
12.4083 | -105.6444
12.412 | -105.6489
12.4156 | -105.6562
12.4193 | -105.6648
12.423 | -105.6732
12.4267 | -105.6849
12.4303 | -105.696
12.434 | -105.7071
12.4377 | -105.721
12.4413 | -105.7353
12.445 | -105.7533
12.4487 | -105.7724
12.4524 | -105.7911
12.456 | -105.8139
12.4597 | -105.8374
12.4634 | -105.8553
12.4671 | -105.8709
12.4707 | -105.886
12.4744 | -105.902
12.4781 | -105.9171
12.4817 | -105.9299
12.4854 | -105.9437
12.4891 | -105.9539
12.4928 | -105.9622
12.4964 | -105.9692
12.5001 | -105.9766
12.5038 | -105.9834
12.5075 | -105.987
12.5111 | -105.9934
12.5148 | -105.9988
12.5185 | -106.0009
12.5222 | -105.999
12.5258 | -105.9952
12.5295 | -105.9954
12.5332 | -105.9955
12.5368 | -105.9932
12.5405 | -105.9908
12.5442 | -105.9821
12.5479 | -105.9715
12.5515 | -105.961
12.5552 | -105.9502
12.5589 | -105.9392
12.5626 | -105.9254
12.5662 | -105.9144
12.5699 | -105.9043
12.5736 | -105.8937
12.5773 | -105.8834
12.5809 | -105.8736
12.5846 | -105.8706
12.5883 | -105.8632
12.5919 | -105.8508
12.5956 | -105.8412
12.5993 | -105.8342
12.603 | -105.8306
12.6066 | -105.8268
12.6103 | -105.8238
12.614 | -105.8191
12.6177 | -105.8133
12.6213 | -105.8077
12.625 | -105.802
12.6287 | -105.7966
12.6324 | -105.7933
12.636 | -105.7892
12.6397 | -105.7839
12.6434 | -105.7844
12.647 | -105.7821
12.6507 | -105.7816
12.6544 | -105.7867
12.6581 | -105.7917
12.6617 | -105.7906
12.6654 | -105.7908
12.6691 | -105.7956
12.6728 | -105.7991
12.6764 | -105.8017
12.6801 | -105.8005
12.6838 | -105.7998
12.6875 | -105.7999
12.6911 | -105.7997
12.6948 | -105.8014
12.6985 | -105.8013
12.7021 | -105.8021
12.7058 | -105.801
12.7095 | -105.7958
12.7132 | -105.7925
12.7168 | -105.7958
12.7205 | -105.7947
12.7242 | -105.7972
12.7279 | -105.8048
12.7315 | -105.8128
12.7352 | -105.8197
12.7389 | -105.8276
12.7425 | -105.8386
12.7462 | -105.8477
12.7499 | -105.8548
12.7536 | -105.8663
12.7572 | -105.8796
12.7609 | -105.8921
12.7646 | -105.9019
12.7683 | -105.9131
12.7719 | -105.9283
12.7756 | -105.9431
12.7793 | -105.955
12.783 | -105.9663
12.7866 | -105.9732
12.7903 | -105.9794
12.794 | -105.9821
12.7976 | -105.9845
12.8013 | -105.9882
12.805 | -105.9897
12.8087 | -105.9909
12.8123 | -105.9911
12.816 | -105.9888
12.8197 | -105.9878
12.8234 | -105.9876
12.827 | -105.9841
12.8307 | -105.986
12.8344 | -105.9885
12.8381 | -105.9854
12.8417 | -105.9837
12.8454 | -105.989
12.8491 | -105.9949
12.8527 | -105.997
12.8564 | -106.0012
12.8601 | -106.0024
12.8638 | -106.0064
12.8674 | -106.0118
12.8711 | -106.016
12.8748 | -106.0216
12.8785 | -106.0261
12.8821 | -106.0311
12.8858 | -106.0359
12.8895 | -106.0401
12.8932 | -106.0443
12.8968 | -106.0483
12.9005 | -106.0522
12.9042 | -106.0539
12.9078 | -106.0547
12.9115 | -106.0551
12.9152 | -106.0546
12.9189 | -106.0558
12.9225 | -106.0573
12.9262 | -106.0583
12.9299 | -106.0552
12.9336 | -106.0567
12.9372 | -106.0551
12.9409 | -106.0505
12.9446 | -106.0495
12.9483 | -106.0495
12.9519 | -106.0513
12.9556 | -106.0505
12.9593 | -106.0513
12.9629 | -106.0533
12.9666 | -106.0553
12.9703 | -106.0613
12.974 | -106.0692
12.9776 | -106.0779
12.9813 | -106.0864
12.985 | -106.0934
12.9887 | -106.0965
12.9923 | -106.1026
12.996 | -106.1073
12.9997 | -106.1132
13.0034 | -106.1205
13.007 | -106.1249
13.0107 | -106.1293
13.0144 | -106.1352
13.018 | -106.1387
13.0217 | -106.1415
13.0254 | -106.1488
13.0291 | -106.1559
13.0327 | -106.1596
13.0364 | -106.1621
13.0401 | -106.1668
13.0438 | -106.1716
13.0474 | -106.1782
13.0511 | -106.1791
13.0548 | -106.1763
13.0585 | -106.1751
13.0621 | -106.1742
13.0658 | -106.1736
13.0695 | -106.1737
13.0731 | -106.1722
13.0768 | -106.1701
13.0805 | -106.1621
13.0842 | -106.155
13.0878 | -106.1481
13.0915 | -106.1436
13.0952 | -106.1393
13.0989 | -106.1337
13.1025 | -106.1285
13.1062 | -106.1204
13.1099 | -106.1108
13.1136 | -106.1066
13.1172 | -106.1065
13.1209 | -106.1034
13.1246 | -106.1039
13.1282 | -106.1048
13.1319 | -106.1018
13.1356 | -106.099
13.1393 | -106.097
13.1429 | -106.0955
13.1466 | -106.0938
13.1503 | -106.0972
13.154 | -106.1018
13.1576 | -106.0999
13.1613 | -106.0995
13.165 | -106.0992
13.1686 | -106.0992
13.1723 | -106.1016
13.176 | -106.1056
13.1797 | -106.1043
13.1833 | -106.105
13.187 | -106.1048
13.1907 | -106.1023
13.1944 | -106.0988
13.198 | -106.0952
13.2017 | -106.0916
13.2054 | -106.0887
13.2091 | -106.0869
13.2127 | -106.0834
13.2164 | -106.0827
13.2201 | -106.0838
13.2237 | -106.085
13.2274 | -106.0877
13.2311 | -106.089
13.2348 | -106.0893
13.2384 | -106.0939
13.2421 | -106.0995
13.2458 | -106.1121
13.2495 | -106.1216
13.2531 | -106.1298
13.2568 | -106.1389
13.2605 | -106.1492
13.2642 | -106.1616
13.2678 | -106.1744
13.2715 | -106.1906
13.2752 | -106.2032
13.2788 | -106.2165
13.2825 | -106.2309
13.2862 | -106.2472
13.2899 | -106.2626
13.2935 | -106.28
13.2972 | -106.2998
13.3009 | -106.319
13.3046 | -106.3356
13.3082 | -106.3504
13.3119 | -106.3662
13.3156 | -106.3802
13.3193 | -106.3923
13.3229 | -106.4052
13.3266 | -106.4173
13.3303 | -106.4279
13.3339 | -106.4348
13.3376 | -106.4421
13.3413 | -106.4434
13.345 | -106.4414
13.3486 | -106.4398
13.3523 | -106.435
13.356 | -106.4289
13.3597 | -106.4232
13.3633 | -106.413
13.367 | -106.3997
13.3707 | -106.3898
13.3744 | -106.3776
13.378 | -106.3657
13.3817 | -106.3561
13.3854 | -106.3462
13.389 | -106.337
13.3927 | -106.3289
13.3964 | -106.3212
13.4001 | -106.315
13.4037 | -106.3125
13.4074 | -106.309
13.4111 | -106.3111
13.4148 | -106.3147
13.4184 | -106.3163
13.4221 | -106.3212
13.4258 | -106.3296
13.4294 | -106.3396
13.4331 | -106.3518
13.4368 | -106.3601
13.4405 | -106.3638
13.4441 | -106.3691
13.4478 | -106.3749
13.4515 | -106.3824
13.4552 | -106.3883
13.4588 | -106.3924
13.4625 | -106.3954
13.4662 | -106.3976
13.4699 | -106.3952
13.4735 | -106.3911
13.4772 | -106.3849
13.4809 | -106.3729
13.4845 | -106.3579
13.4882 | -106.3407
13.4919 | -106.3248
13.4956 | -106.3093
13.4992 | -106.2896
13.5029 | -106.2703
13.5066 | -106.2495
13.5103 | -106.2259
13.5139 | -106.2022
13.5176 | -106.1818
13.5213 | -106.1628
13.525 | -106.1429
13.5286 | -106.1249
13.5323 | -106.1066
13.536 | -106.0896
13.5396 | -106.075
13.5433 | -106.0606
13.547 | -106.0448
13.5507 | -106.0301
13.5543 | -106.0181
13.558 | -106.0068
13.5617 | -105.9973
13.5654 | -105.9892
13.569 | -105.9792
13.5727 | -105.9715
13.5764 | -105.9653
13.5801 | -105.9611
13.5837 | -105.9579
13.5874 | -105.9538
13.5911 | -105.9509
13.5947 | -105.95
13.5984 | -105.9481
13.6021 | -105.9466
13.6058 | -105.9455
13.6094 | -105.9453
13.6131 | -105.9477
13.6168 | -105.9495
13.6205 | -105.9519
13.6241 | -105.9535
13.6278 | -105.9537
13.6315 | -105.9563
13.6352 | -105.9613
13.6388 | -105.9665
13.6425 | -105.9689
13.6462 | -105.9704
13.6498 | -105.9754
13.6535 | -105.9812
13.6572 | -105.9877
13.6609 | -105.9902
13.6645 | -105.994
13.6682 | -105.9977
13.6719 | -105.999
13.6756 | -106.0026
13.6792 | -106.0062
13.6829 | -106.0073
13.6866 | -106.0068
13.6903 | -106.0041
13.6939 | -106.0016
13.6976 | -105.9978
13.7013 | -105.9957
13.7049 | -105.9946
13.7086 | -105.9926
13.7123 | -105.99
13.716 | -105.9895
13.7196 | -105.9908
13.7233 | -105.9898
13.727 | -105.992
13.7307 | -105.9967
13.7343 | -106.0039
13.738 | -106.0123
13.7417 | -106.0247
13.7454 | -106.0364
13.749 | -106.0453
13.7527 | -106.0565
13.7564 | -106.0711
13.76 | -106.0833
13.7637 | -106.0987
13.7674 | -106.1129
13.7711 | -106.1276
13.7747 | -106.138
13.7784 | -106.1487
13.7821 | -106.1581
13.7858 | -106.1676
13.7894 | -106.1774
13.7931 | -106.1841
13.7968 | -106.1922
13.8004 | -106.201
13.8041 | -106.2061
13.8078 | -106.2094
13.8115 | -106.2097
13.8151 | -106.2118
13.8188 | -106.2162
13.8225 | -106.2182
13.8262 | -106.2201
13.8298 | -106.2229
13.8335 | -106.2272
13.8372 | -106.2311
13.8409 | -106.2353
13.8445 | -106.2395
13.8482 | -106.2406
13.8519 | -106.24
13.8555 | -106.242
13.8592 | -106.2486
13.8629 | -106.2561
13.8666 | -106.2622
13.8702 | -106.2692
13.8739 | -106.2742
13.8776 | -106.2782
13.8813 | -106.2812
13.8849 | -106.2805
13.8886 | -106.2796
13.8923 | -106.2808
13.896 | -106.2801
13.8996 | -106.2765
13.9033 | -106.2743
13.907 | -106.2719
13.9106 | -106.2652
13.9143 | -106.2568
13.918 | -106.2493
13.9217 | -106.2437
13.9253 | -106.2345
13.929 | -106.2252
13.9327 | -106.2183
13.9364 | -106.2143
13.94 | -106.2097
13.9437 | -106.2053
13.9474 | -106.2033
13.9511 | -106.2012
13.9547 | -106.1957
13.9584 | -106.1805
13.9621 | -106.1589
13.9657 | -106.1422
13.9694 | -106.1272
13.9731 | -106.1149
13.9768 | -106.1037
13.9804 | -106.0908
13.9841 | -106.0753
13.9878 | -106.0554
13.9915 | -106.0327
13.9951 | -106.0071
13.9988 | -105.978
14.0025 | -105.9478
14.0062 | -105.9183
14.0098 | -105.8895
14.0135 | -105.8568
14.0172 | -105.8194
14.0208 | -105.7796
14.0245 | -105.741
14.0282 | -105.7033
14.0319 | -105.6649
14.0355 | -105.6284
14.0392 | -105.5952
14.0429 | -105.5646
14.0466 | -105.5374
14.0502 | -105.5162
14.0539 | -105.4996
14.0576 | -105.4854
14.0613 | -105.4764
14.0649 | -105.4715
14.0686 | -105.4732
14.0723 | -105.4801
14.0759 | -105.4927
14.0796 | -105.5111
14.0833 | -105.5345
14.087 | -105.5645
14.0906 | -105.5995
14.0943 | -105.6372
14.098 | -105.6763
14.1017 | -105.7173
14.1053 | -105.76
14.109 | -105.8064
14.1127 | -105.8571
14.1164 | -105.9069
14.12 | -105.9534
14.1237 | -106.0025
14.1274 | -106.0514
14.131 | -106.0978
14.1347 | -106.1398
14.1384 | -106.1794
14.1421 | -106.2174
14.1457 | -106.2537
14.1494 | -106.2871
14.1531 | -106.3169
14.1568 | -106.344
14.1604 | -106.3775
14.1641 | -106.4073
14.1678 | -106.4269
14.1715 | -106.4387
14.1751 | -106.444
14.1788 | -106.4455
14.1825 | -106.4521
14.1861 | -106.4553
14.1898 | -106.457
14.1935 | -106.4589
14.1972 | -106.4605
14.2008 | -106.4589
14.2045 | -106.4526
14.2082 | -106.4469
14.2119 | -106.4401
14.2155 | -106.4364
14.2192 | -106.4332
14.2229 | -106.4262
14.2265 | -106.4192
14.2302 | -106.4131
14.2339 | -106.405
14.2376 | -106.3946
14.2412 | -106.3812
14.2449 | -106.3681
14.2486 | -106.353
14.2523 | -106.3363
14.2559 | -106.3199
14.2596 | -106.3014
14.2633 | -106.2832
14.267 | -106.2668
14.2706 | -106.2486
14.2743 | -106.2323
14.278 | -106.2169
14.2816 | -106.1987
14.2853 | -106.1787
14.289 | -106.1608
14.2927 | -106.1436
14.2963 | -106.1278
14.3 | -106.1158
14.3037 | -106.1048
14.3074 | -106.0995
14.311 | -106.0925
14.3147 | -106.0859
14.3184 | -106.0787
14.3221 | -106.0767
14.3257 | -106.0783
14.3294 | -106.0804
14.3331 | -106.0835
14.3367 | -106.0897
14.3404 | -106.0967
14.3441 | -106.1069
14.3478 | -106.1185
14.3514 | -106.133
14.3551 | -106.1498
14.3588 | -106.1686
14.3625 | -106.191
14.3661 | -106.2142
14.3698 | -106.2378
14.3735 | -106.259
14.3772 | -106.2833
14.3808 | -106.308
14.3845 | -106.3344
14.3882 | -106.3571
14.3918 | -106.3814
14.3955 | -106.4072
14.3992 | -106.4302
14.4029 | -106.4515
14.4065 | -106.4753
14.4102 | -106.4979
14.4139 | -106.5143
14.4176 | -106.5277
14.4212 | -106.5403
14.4249 | -106.5507
14.4286 | -106.5586
14.4323 | -106.5629
14.4359 | -106.5649
14.4396 | -106.5649
14.4433 | -106.5615
14.4469 | -106.556
14.4506 | -106.5457
14.4543 | -106.532
14.458 | -106.5196
14.4616 | -106.5053
14.4653 | -106.4894
14.469 | -106.4712
14.4727 | -106.4527
14.4763 | -106.4349
14.48 | -106.4156
14.4837 | -106.3979
14.4873 | -106.3827
14.491 | -106.3696
14.4947 | -106.3523
14.4984 | -106.3357
14.502 | -106.3208
14.5057 | -106.3104
14.5094 | -106.3005
14.5131 | -106.2891
14.5167 | -106.2787
14.5204 | -106.2681
14.5241 | -106.2598
14.5278 | -106.2516
14.5314 | -106.2435
14.5351 | -106.2378
14.5388 | -106.2319
14.5424 | -106.2284
14.5461 | -106.2253
14.5498 | -106.2242
14.5535 | -106.2195
14.5571 | -106.2159
14.5608 | -106.2145
14.5645 | -106.2124
14.5682 | -106.2133
14.5718 | -106.2128
14.5755 | -106.2122
14.5792 | -106.2146
14.5829 | -106.2174
14.5865 | -106.2176
14.5902 | -106.2191
14.5939 | -106.2223
14.5975 | -106.2218
14.6012 | -106.2239
14.6049 | -106.2304
14.6086 | -106.2369
14.6122 | -106.2456
14.6159 | -106.2512
14.6196 | -106.2572
14.6233 | -106.2656
14.6269 | -106.2788
14.6306 | -106.2917
14.6343 | -106.3037
14.638 | -106.318
14.6416 | -106.336
14.6453 | -106.3541
14.649 | -106.3747
14.6526 | -106.3966
14.6563 | -106.4199
14.66 | -106.4446
14.6637 | -106.4702
14.6673 | -106.4916
14.671 | -106.5102
14.6747 | -106.5295
14.6784 | -106.5482
14.682 | -106.5644
14.6857 | -106.5794
14.6894 | -106.5933
14.6931 | -106.6053
14.6967 | -106.6145
14.7004 | -106.6209
14.7041 | -106.6291
14.7077 | -106.6334
14.7114 | -106.6366
14.7151 | -106.635
14.7188 | -106.6338
14.7224 | -106.6336
14.7261 | -106.6294
14.7298 | -106.6237
14.7335 | -106.6159
14.7371 | -106.6108
14.7408 | -106.6029
14.7445 | -106.5934
14.7482 | -106.5846
14.7518 | -106.5749
14.7555 | -106.5638
14.7592 | -106.5518
14.7628 | -106.544
14.7665 | -106.5377
14.7702 | -106.5257
14.7739 | -106.5128
14.7775 | -106.4975
14.7812 | -106.4829
14.7849 | -106.473
14.7886 | -106.462
14.7922 | -106.4521
14.7959 | -106.4421
14.7996 | -106.4315
14.8033 | -106.4212
14.8069 | -106.4148
14.8106 | -106.4097
14.8143 | -106.4006
14.8179 | -106.3914
14.8216 | -106.386
14.8253 | -106.3832
14.829 | -106.3825
14.8326 | -106.3792
14.8363 | -106.3789
14.84 | -106.3761
14.8437 | -106.3738
14.8473 | -106.3718
14.851 | -106.3732
14.8547 | -106.3728
14.8583 | -106.3724
14.862 | -106.3707
14.8657 | -106.3704
14.8694 | -106.3712
14.873 | -106.371
14.8767 | -106.3693
14.8804 | -106.3696
14.8841 | -106.3722
14.8877 | -106.3739
14.8914 | -106.3735
14.8951 | -106.3743
14.8988 | -106.3777
14.9024 | -106.379
14.9061 | -106.379
14.9098 | -106.3803
14.9134 | -106.3827
14.9171 | -106.3848
14.9208 | -106.3897
14.9245 | -106.3905
14.9281 | -106.3897
14.9318 | -106.3884
14.9355 | -106.3891
14.9392 | -106.3884
14.9428 | -106.3898
14.9465 | -106.3896
14.9502 | -106.3887
14.9539 | -106.3884
14.9575 | -106.3861
14.9612 | -106.3839
14.9649 | -106.383
14.9685 | -106.3784
14.9722 | -106.3741
14.9759 | -106.3694
14.9796 | -106.363
14.9832 | -106.3606
14.9869 | -106.3552
14.9906 | -106.3518
14.9943 | -106.3472
14.9979 | -106.3366
15.0016 | -106.3268
15.0053 | -106.3198
15.009 | -106.3114
15.0126 | -106.3044
15.0163 | -106.296
15.02 | -106.286
15.0236 | -106.2759
15.0273 | -106.2669
15.031 | -106.2582
15.0347 | -106.2487
15.0383 | -106.2398
15.042 | -106.2309
15.0457 | -106.2226
15.0494 | -106.2152
15.053 | -106.2087
15.0567 | -106.2036
15.0604 | -106.2009
15.0641 | -106.2013
15.0677 | -106.2011
15.0714 | -106.2033
15.0751 | -106.2083
15.0787 | -106.2156
15.0824 | -106.2218
15.0861 | -106.2274
15.0898 | -106.2375
15.0934 | -106.251
15.0971 | -106.2614
15.1008 | -106.2736
15.1045 | -106.2867
15.1081 | -106.3009
15.1118 | -106.3163
15.1155 | -106.3311
15.1192 | -106.348
15.1228 | -106.3666
15.1265 | -106.3864
15.1302 | -106.4055
15.1338 | -106.4249
15.1375 | -106.4419
15.1412 | -106.4626
15.1449 | -106.4867
15.1485 | -106.5124
15.1522 | -106.5362
15.1559 | -106.5623
15.1596 | -106.5834
15.1632 | -106.6041
15.1669 | -106.6245
15.1706 | -106.6433
15.1743 | -106.6621
15.1779 | -106.6781
15.1816 | -106.6949
15.1853 | -106.7132
15.1889 | -106.7289
15.1926 | -106.742
15.1963 | -106.7538
15.2 | -106.7641
15.2036 | -106.7756
15.2073 | -106.7793
15.211 | -106.779
15.2147 | -106.7793
15.2183 | -106.7757
15.222 | -106.7682
15.2257 | -106.7617
15.2293 | -106.7521
15.233 | -106.7433
15.2367 | -106.7323
15.2404 | -106.7176
15.244 | -106.7029
15.2477 | -106.6876
15.2514 | -106.6667
15.2551 | -106.6479
15.2587 | -106.6285
15.2624 | -106.6099
15.2661 | -106.5878
15.2698 | -106.5644
15.2734 | -106.5394
15.2771 | -106.5148
15.2808 | -106.4899
15.2844 | -106.4628
15.2881 | -106.4347
15.2918 | -106.4078
15.2955 | -106.3846
15.2991 | -106.3641
15.3028 | -106.3453
15.3065 | -106.3259
15.3102 | -106.3076
15.3138 | -106.2912
15.3175 | -106.2779
15.3212 | -106.2658
15.3249 | -106.2543
15.3285 | -106.2449
15.3322 | -106.2381
15.3359 | -106.2332
15.3395 | -106.2301
15.3432 | -106.2284
15.3469 | -106.2286
15.3506 | -106.231
15.3542 | -106.2368
15.3579 | -106.2424
15.3616 | -106.2484
15.3653 | -106.2593
15.3689 | -106.2658
15.3726 | -106.2714
15.3763 | -106.2804
15.38 | -106.2925
15.3836 | -106.306
15.3873 | -106.3185
15.391 | -106.3314
15.3946 | -106.3454
15.3983 | -106.3564
15.402 | -106.364
15.4057 | -106.3729
15.4093 | -106.3818
15.413 | -106.3883
15.4167 | -106.3956
15.4204 | -106.4042
15.424 | -106.4121
15.4277 | -106.4217
15.4314 | -106.4318
15.4351 | -106.4377
15.4387 | -106.4415
15.4424 | -106.4431
15.4461 | -106.445
15.4497 | -106.4469
15.4534 | -106.4461
15.4571 | -106.4449
15.4608 | -106.4408
15.4644 | -106.4364
15.4681 | -106.4303
15.4718 | -106.424
15.4755 | -106.4165
15.4791 | -106.4087
15.4828 | -106.3997
15.4865 | -106.389
15.4902 | -106.3789
15.4938 | -106.3707
15.4975 | -106.3605
15.5012 | -106.348
15.5048 | -106.3376
15.5085 | -106.3289
15.5122 | -106.3211
15.5159 | -106.3098
15.5195 | -106.2993
15.5232 | -106.2908
15.5269 | -106.2821
15.5306 | -106.2737
15.5342 | -106.2664
15.5379 | -106.2568
15.5416 | -106.252
15.5452 | -106.2493
15.5489 | -106.2466
15.5526 | -106.2452
15.5563 | -106.2452
15.5599 | -106.2432
15.5636 | -106.2431
15.5673 | -106.2461
15.571 | -106.2499
15.5746 | -106.2558
15.5783 | -106.2607
15.582 | -106.2652
15.5857 | -106.2727
15.5893 | -106.2763
15.593 | -106.2809
15.5967 | -106.2891
15.6003 | -106.2931
15.604 | -106.2966
15.6077 | -106.3038
15.6114 | -106.3117
15.615 | -106.3195
15.6187 | -106.326
15.6224 | -106.3349
15.6261 | -106.3439
15.6297 | -106.3517
15.6334 | -106.3583
15.6371 | -106.3657
15.6408 | -106.3707
15.6444 | -106.3722
15.6481 | -106.3721
15.6518 | -106.3724
15.6554 | -106.3727
15.6591 | -106.3738
15.6628 | -106.373
15.6665 | -106.3708
15.6701 | -106.3674
15.6738 | -106.361
15.6775 | -106.3568
15.6812 | -106.3533
15.6848 | -106.3504
15.6885 | -106.3433
15.6922 | -106.3345
15.6959 | -106.3272
15.6995 | -106.3188
15.7032 | -106.3109
15.7069 | -106.3021
15.7105 | -106.293
15.7142 | -106.283
15.7179 | -106.2723
15.7216 | -106.2628
15.7252 | -106.2572
15.7289 | -106.2505
15.7326 | -106.2419
15.7363 | -106.2341
15.7399 | -106.2296
15.7436 | -106.225
15.7473 | -106.2197
15.751 | -106.2145
15.7546 | -106.2092
15.7583 | -106.2054
15.762 | -106.2024
15.7656 | -106.1962
15.7693 | -106.1916
15.773 | -106.1886
15.7767 | -106.185
15.7803 | -106.1828
15.784 | -106.1822
15.7877 | -106.1804
15.7914 | -106.179
15.795 | -106.1755
15.7987 | -106.1728
15.8024 | -106.1741
15.8061 | -106.1758
15.8097 | -106.1741
15.8134 | -106.1753
15.8171 | -106.1758
15.8207 | -106.1756
15.8244 | -106.1762
15.8281 | -106.1766
15.8318 | -106.1703
15.8354 | -106.1668
15.8391 | -106.1639
15.8428 | -106.161
15.8465 | -106.156
15.8501 | -106.1521
15.8538 | -106.1507
15.8575 | -106.1477
15.8612 | -106.1437
15.8648 | -106.1416
15.8685 | -106.1403
15.8722 | -106.1403
15.8758 | -106.141
15.8795 | -106.1402
15.8832 | -106.1412
15.8869 | -106.145
15.8905 | -106.1483
15.8942 | -106.1495
15.8979 | -106.1573
15.9016 | -106.1622
15.9052 | -106.168
15.9089 | -106.1744
15.9126 | -106.1811
15.9162 | -106.1892
15.9199 | -106.1985
15.9236 | -106.2075
15.9273 | -106.2158
15.9309 | -106.2244
15.9346 | -106.2355
15.9383 | -106.2492
15.942 | -106.2616
15.9456 | -106.2724
15.9493 | -106.2839
15.953 | -106.295
15.9567 | -106.3051
15.9603 | -106.3172
15.964 | -106.3288
15.9677 | -106.3401
15.9713 | -106.3531
15.975 | -106.3667
15.9787 | -106.3793
15.9824 | -106.3928
15.986 | -106.4083
15.9897 | -106.4216
15.9934 | -106.4339
15.9971 | -106.4451
16.0007 | -106.4559
16.0044 | -106.4658
16.0081 | -106.4753
16.0118 | -106.4845
16.0154 | -106.4936
16.0191 | -106.503
16.0228 | -106.5144
16.0264 | -106.5245
16.0301 | -106.5323
16.0338 | -106.54
16.0375 | -106.55
16.0411 | -106.5603
16.0448 | -106.5699
16.0485 | -106.578
16.0522 | -106.5861
16.0558 | -106.5953
16.0595 | -106.6078
16.0632 | -106.6169
16.0669 | -106.624
16.0705 | -106.6335
16.0742 | -106.6447
16.0779 | -106.6537
16.0815 | -106.6594
16.0852 | -106.6683
16.0889 | -106.6747
16.0926 | -106.6806
16.0962 | -106.6858
16.0999 | -106.6903
16.1036 | -106.6928
16.1073 | -106.694
16.1109 | -106.6938
16.1146 | -106.6919
16.1183 | -106.692
16.122 | -106.6922
16.1256 | -106.6913
16.1293 | -106.6885
16.133 | -106.6862
16.1366 | -106.6814
16.1403 | -106.6778
16.144 | -106.6754
16.1477 | -106.6727
16.1513 | -106.6713
16.155 | -106.6692
16.1587 | -106.6689
16.1624 | -106.6706
16.166 | -106.6715
16.1697 | -106.6698
16.1734 | -106.6717
16.1771 | -106.6748
16.1807 | -106.6767
16.1844 | -106.6786
16.1881 | -106.6836
16.1917 | -106.6875
16.1954 | -106.6944
16.1991 | -106.6995
16.2028 | -106.7025
16.2064 | -106.7092
16.2101 | -106.7205
16.2138 | -106.7283
16.2175 | -106.7333
16.2211 | -106.7394
16.2248 | -106.7455
16.2285 | -106.7501
16.2322 | -106.7525
16.2358 | -106.7569
16.2395 | -106.7594
16.2432 | -106.7627
16.2468 | -106.7633
16.2505 | -106.7591
16.2542 | -106.7556
16.2579 | -106.7515
16.2615 | -106.7486
16.2652 | -106.745
16.2689 | -106.7376
16.2726 | -106.7286
16.2762 | -106.7182
16.2799 | -106.7098
16.2836 | -106.7044
16.2872 | -106.6969
16.2909 | -106.6906
16.2946 | -106.6849
16.2983 | -106.6787
16.3019 | -106.6686
16.3056 | -106.6624
16.3093 | -106.6568
16.313 | -106.6523
16.3166 | -106.6466
16.3203 | -106.6437
16.324 | -106.6396
16.3277 | -106.6362
16.3313 | -106.6363
16.335 | -106.6356
16.3387 | -106.638
16.3423 | -106.6409
16.346 | -106.6434
16.3497 | -106.6439
16.3534 | -106.6432
16.357 | -106.646
16.3607 | -106.6483
16.3644 | -106.6543
16.3681 | -106.6601
16.3717 | -106.6624
16.3754 | -106.6644
16.3791 | -106.67
16.3828 | -106.6738
16.3864 | -106.6766
16.3901 | -106.6789
16.3938 | -106.6791
16.3974 | -106.6808
16.4011 | -106.6852
16.4048 | -106.6866
16.4085 | -106.6867
16.4121 | -106.6879
16.4158 | -106.69
16.4195 | -106.6907
16.4232 | -106.6904
16.4268 | -106.6848
16.4305 | -106.6802
16.4342 | -106.6772
16.4379 | -106.6738
16.4415 | -106.6686
16.4452 | -106.661
16.4489 | -106.6555
16.4525 | -106.6527
16.4562 | -106.6509
16.4599 | -106.6491
16.4636 | -106.6444
16.4672 | -106.6383
16.4709 | -106.6322
16.4746 | -106.626
16.4783 | -106.6171
16.4819 | -106.6115
16.4856 | -106.6113
16.4893 | -106.611
16.493 | -106.6088
16.4966 | -106.6067
16.5003 | -106.6071
16.504 | -106.6107
16.5076 | -106.615
16.5113 | -106.6195
16.515 | -106.624
16.5187 | -106.6277
16.5223 | -106.6334
16.526 | -106.64
16.5297 | -106.6473
16.5334 | -106.6573
16.537 | -106.6693
16.5407 | -106.6794
16.5444 | -106.6905
16.5481 | -106.7026
16.5517 | -106.7143
16.5554 | -106.7244
16.5591 | -106.7371
16.5627 | -106.7507
16.5664 | -106.7618
16.5701 | -106.7728
16.5738 | -106.784
16.5774 | -106.7942
16.5811 | -106.8038
16.5848 | -106.8147
16.5885 | -106.8263
16.5921 | -106.8349
16.5958 | -106.8426
16.5995 | -106.8457
16.6031 | -106.8503
16.6068 | -106.857
16.6105 | -106.8614
16.6142 | -106.8618
16.6178 | -106.8629
16.6215 | -106.8641
16.6252 | -106.8641
16.6289 | -106.8628
16.6325 | -106.86
16.6362 | -106.856
16.6399 | -106.8533
16.6436 | -106.8481
16.6472 | -106.8423
16.6509 | -106.8355
16.6546 | -106.8273
16.6582 | -106.8179
16.6619 | -106.8113
16.6656 | -106.8038
16.6693 | -106.7974
16.6729 | -106.7907
16.6766 | -106.7855
16.6803 | -106.7788
16.684 | -106.7719
16.6876 | -106.7666
16.6913 | -106.7595
16.695 | -106.7513
16.6987 | -106.7449
16.7023 | -106.7434
16.706 | -106.7429
16.7097 | -106.7402
16.7133 | -106.7401
16.717 | -106.7426
16.7207 | -106.7427
16.7244 | -106.74
16.728 | -106.7366
16.7317 | -106.7353
16.7354 | -106.7349
16.7391 | -106.7357
16.7427 | -106.7387
16.7464 | -106.7423
16.7501 | -106.7477
16.7538 | -106.7523
16.7574 | -106.7554
16.7611 | -106.7601
16.7648 | -106.7654
16.7684 | -106.7706
16.7721 | -106.7773
16.7758 | -106.783
16.7795 | -106.7875
16.7831 | -106.7949
16.7868 | -106.8024
16.7905 | -106.8108
16.7942 | -106.8219
16.7978 | -106.8313
16.8015 | -106.839
16.8052 | -106.848
16.8089 | -106.8553
16.8125 | -106.8624
16.8162 | -106.8695
16.8199 | -106.8778
16.8235 | -106.8854
16.8272 | -106.8942
16.8309 | -106.9029
16.8346 | -106.9112
16.8382 | -106.919
16.8419 | -106.9261
16.8456 | -106.933
16.8493 | -106.9406
16.8529 | -106.9484
16.8566 | -106.9569
16.8603 | -106.9623
16.864 | -106.9676
16.8676 | -106.9753
16.8713 | -106.9843
16.875 | -106.9908
16.8786 | -106.9971
16.8823 | -107.0033
16.886 | -107.0095
16.8897 | -107.0165
16.8933 | -107.0222
16.897 | -107.0287
16.9007 | -107.0319
16.9044 | -107.0363
16.908 | -107.0402
16.9117 | -107.0447
16.9154 | -107.0487
16.919 | -107.0527
16.9227 | -107.0588
16.9264 | -107.0618
16.9301 | -107.0662
16.9337 | -107.0709
16.9374 | -107.0742
16.9411 | -107.0735
16.9448 | -107.0729
16.9484 | -107.0747
16.9521 | -107.073
16.9558 | -107.0714
16.9595 | -107.073
16.9631 | -107.0746
16.9668 | -107.0781
16.9705 | -107.0805
16.9741 | -107.0853
16.9778 | -107.0881
16.9815 | -107.0902
16.9852 | -107.0923
16.9888 | -107.0924
16.9925 | -107.0946
16.9962 | -107.0998
16.9999 | -107.1063
17.0035 | -107.1105
17.0072 | -107.1139
17.0109 | -107.1194
17.0146 | -107.125
17.0182 | -107.1301
17.0219 | -107.1362
17.0256 | -107.1394
17.0292 | -107.1413
17.0329 | -107.1454
17.0366 | -107.1476
17.0403 | -107.1513
17.0439 | -107.1579
17.0476 | -107.1633
17.0513 | -107.1702
17.055 | -107.1731
17.0586 | -107.1766
17.0623 | -107.1819
17.066 | -107.1855
17.0697 | -107.1873
17.0733 | -107.1907
17.077 | -107.1963
17.0807 | -107.2032
17.0843 | -107.21
17.088 | -107.2165
17.0917 | -107.2221
17.0954 | -107.2233
17.099 | -107.2241
17.1027 | -107.2284
17.1064 | -107.2324
17.1101 | -107.2347
17.1137 | -107.2335
17.1174 | -107.2346
17.1211 | -107.2371
17.1248 | -107.2395
17.1284 | -107.241
17.1321 | -107.2394
17.1358 | -107.2395
17.1394 | -107.2403
17.1431 | -107.2396
17.1468 | -107.2372
17.1505 | -107.2342
17.1541 | -107.2293
17.1578 | -107.2231
17.1615 | -107.2177
17.1652 | -107.2103
17.1688 | -107.2039
17.1725 | -107.1974
17.1762 | -107.1925
17.1799 | -107.1883
17.1835 | -107.1849
17.1872 | -107.1818
17.1909 | -107.1756
17.1945 | -107.1705
17.1982 | -107.1635
17.2019 | -107.1514
17.2056 | -107.1371
17.2092 | -107.1238
17.2129 | -107.1101
17.2166 | -107.0999
17.2203 | -107.0877
17.2239 | -107.0759
17.2276 | -107.0639
17.2313 | -107.0505
17.235 | -107.039
17.2386 | -107.0265
17.2423 | -107.0112
17.246 | -106.9962
17.2496 | -106.9837
17.2533 | -106.9705
17.257 | -106.9567
17.2607 | -106.9418
17.2643 | -106.9322
17.268 | -106.9237
17.2717 | -106.9164
17.2754 | -106.9111
17.279 | -106.9043
17.2827 | -106.9005
17.2864 | -106.896
17.2901 | -106.8908
17.2937 | -106.889
17.2974 | -106.89
17.3011 | -106.8952
17.3047 | -106.8997
17.3084 | -106.9036
17.3121 | -106.9101
17.3158 | -106.918
17.3194 | -106.9247
17.3231 | -106.9315
17.3268 | -106.9364
17.3305 | -106.9423
17.3341 | -106.9491
17.3378 | -106.957
17.3415 | -106.9689
17.3451 | -106.9784
17.3488 | -106.9885
17.3525 | -106.9992
17.3562 | -107.0082
17.3598 | -107.0203
17.3635 | -107.0296
17.3672 | -107.0372
17.3709 | -107.0451
17.3745 | -107.0515
17.3782 | -107.0571
17.3819 | -107.0603
17.3856 | -107.0632
17.3892 | -107.0691
17.3929 | -107.0725
17.3966 | -107.076
17.4002 | -107.0809
17.4039 | -107.0849
17.4076 | -107.0862
17.4113 | -107.0842
17.4149 | -107.0797
17.4186 | -107.0753
17.4223 | -107.0701
17.426 | -107.0631
17.4296 | -107.0571
17.4333 | -107.0499
17.437 | -107.0425
17.4407 | -107.035
17.4443 | -107.0273
17.448 | -107.0197
17.4517 | -107.0127
17.4553 | -107.0031
17.459 | -106.9907
17.4627 | -106.979
17.4664 | -106.9698
17.47 | -106.9624
17.4737 | -106.9535
17.4774 | -106.9468
17.4811 | -106.9411
17.4847 | -106.9388
17.4884 | -106.9342
17.4921 | -106.9323
17.4958 | -106.9306
17.4994 | -106.9291
17.5031 | -106.9295
17.5068 | -106.9283
17.5104 | -106.9287
17.5141 | -106.9273
17.5178 | -106.9272
17.5215 | -106.9272
17.5251 | -106.928
17.5288 | -106.9292
17.5325 | -106.9282
17.5362 | -106.9282
17.5398 | -106.9297
17.5435 | -106.9287
17.5472 | -106.9293
17.5509 | -106.9297
17.5545 | -106.9302
17.5582 | -106.9296
17.5619 | -106.9293
17.5655 | -106.9293
17.5692 | -106.931
17.5729 | -106.9323
17.5766 | -106.9351
17.5802 | -106.9366
17.5839 | -106.9383
17.5876 | -106.9415
17.5913 | -106.9411
17.5949 | -106.9419
17.5986 | -106.9412
17.6023 | -106.9389
17.606 | -106.9393
17.6096 | -106.9392
17.6133 | -106.9406
17.617 | -106.9415
17.6206 | -106.9439
17.6243 | -106.9463
17.628 | -106.9487
17.6317 | -106.9516
17.6353 | -106.9504
17.639 | -106.9468
17.6427 | -106.9452
17.6464 | -106.9444
17.65 | -106.941
17.6537 | -106.9386
17.6574 | -106.9353
17.661 | -106.9347
17.6647 | -106.9353
17.6684 | -106.9312
17.6721 | -106.9272
17.6757 | -106.9255
17.6794 | -106.9242
17.6831 | -106.9241
17.6868 | -106.923
17.6904 | -106.9226
17.6941 | -106.9238
17.6978 | -106.9225
17.7015 | -106.9205
17.7051 | -106.9186
17.7088 | -106.9165
17.7125 | -106.9135
17.7161 | -106.9104
17.7198 | -106.907
17.7235 | -106.9036
17.7272 | -106.8999
17.7308 | -106.8989
17.7345 | -106.8975
17.7382 | -106.8938
17.7419 | -106.8895
17.7455 | -106.8843
17.7492 | -106.879
17.7529 | -106.8734
17.7566 | -106.8696
17.7602 | -106.8675
17.7639 | -106.8633
17.7676 | -106.855
17.7712 | -106.8464
17.7749 | -106.8377
17.7786 | -106.83
17.7823 | -106.8232
17.7859 | -106.8192
17.7896 | -106.8142
17.7933 | -106.808
17.797 | -106.8019
17.8006 | -106.7943
17.8043 | -106.7869
17.808 | -106.7796
17.8117 | -106.7728
17.8153 | -106.7643
17.819 | -106.7557
17.8227 | -106.747
17.8263 | -106.7402
17.83 | -106.7335
17.8337 | -106.7272
17.8374 | -106.7203
17.841 | -106.7134
17.8447 | -106.7075
17.8484 | -106.7013
17.8521 | -106.6949
17.8557 | -106.6918
17.8594 | -106.6866
17.8631 | -106.681
17.8668 | -106.6767
17.8704 | -106.6709
17.8741 | -106.6665
17.8778 | -106.6643
17.8814 | -106.6636
17.8851 | -106.6643
17.8888 | -106.6668
17.8925 | -106.6687
17.8961 | -106.6706
17.8998 | -106.673
17.9035 | -106.6752
17.9072 | -106.6774
17.9108 | -106.6834
17.9145 | -106.6897
17.9182 | -106.6959
17.9219 | -106.7057
17.9255 | -106.7144
17.9292 | -106.7218
17.9329 | -106.7285
17.9365 | -106.7359
17.9402 | -106.7476
17.9439 | -106.7574
17.9476 | -106.7656
17.9512 | -106.7785
17.9549 | -106.7947
17.9586 | -106.8114
17.9623 | -106.8283
17.9659 | -106.8422
17.9696 | -106.8547
17.9733 | -106.8678
17.9769 | -106.8803
17.9806 | -106.8915
17.9843 | -106.9022
17.988 | -106.9139
17.9916 | -106.9257
17.9953 | -106.9365
17.999 | -106.9478
18.0027 | -106.9617
18.0063 | -106.9719
18.01 | -106.9764
18.0137 | -106.9807
18.0174 | -106.9808
18.021 | -106.9868
18.0247 | -106.9956
18.0284 | -107.0022
18.032 | -107.0077
18.0357 | -107.0099
18.0394 | -107.0113
18.0431 | -107.0107
18.0467 | -107.0089
18.0504 | -107.0054
18.0541 | -107.0036
18.0578 | -106.9993
18.0614 | -106.9918
18.0651 | -106.9849
18.0688 | -106.9762
18.0725 | -106.9704
18.0761 | -106.9645
18.0798 | -106.9577
18.0835 | -106.9495
18.0871 | -106.9409
18.0908 | -106.9303
18.0945 | -106.9199
18.0982 | -106.9089
18.1018 | -106.8955
18.1055 | -106.8847
18.1092 | -106.8749
18.1129 | -106.8623
18.1165 | -106.8507
18.1202 | -106.8407
18.1239 | -106.829
18.1276 | -106.821
18.1312 | -106.8122
18.1349 | -106.8021
18.1386 | -106.7929
18.1422 | -106.7841
18.1459 | -106.7757
18.1496 | -106.7682
18.1533 | -106.7618
18.1569 | -106.7542
18.1606 | -106.7476
18.1643 | -106.7442
18.168 | -106.7389
18.1716 | -106.7313
18.1753 | -106.7231
18.179 | -106.7157
18.1827 | -106.7078
18.1863 | -106.7029
18.19 | -106.6981
18.1937 | -106.6925
18.1973 | -106.6851
18.201 | -106.6794
18.2047 | -106.676
18.2084 | -106.6707
18.212 | -106.6611
18.2157 | -106.6519
18.2194 | -106.6458
18.2231 | -106.64
18.2267 | -106.6347
18.2304 | -106.6275
18.2341 | -106.6218
18.2378 | -106.6165
18.2414 | -106.6083
18.2451 | -106.599
18.2488 | -106.5888
18.2524 | -106.5766
18.2561 | -106.5654
18.2598 | -106.5538
18.2635 | -106.5428
18.2671 | -106.5348
18.2708 | -106.5263
18.2745 | -106.5147
18.2782 | -106.5056
18.2818 | -106.4945
18.2855 | -106.4815
18.2892 | -106.4664
18.2929 | -106.4537
18.2965 | -106.4442
18.3002 | -106.433
18.3039 | -106.4209
18.3075 | -106.411
18.3112 | -106.4058
18.3149 | -106.3982
18.3186 | -106.3918
18.3222 | -106.3872
18.3259 | -106.384
18.3296 | -106.3824
18.3333 | -106.3783
18.3369 | -106.376
18.3406 | -106.3728
18.3443 | -106.3689
18.348 | -106.3672
18.3516 | -106.3673
18.3553 | -106.3669
18.359 | -106.3663
18.3626 | -106.3692
18.3663 | -106.371
18.37 | -106.3744
18.3737 | -106.3795
18.3773 | -106.3844
18.381 | -106.3876
18.3847 | -106.3917
18.3884 | -106.3931
18.392 | -106.394
18.3957 | -106.3964
18.3994 | -106.3991
18.403 | -106.403
18.4067 | -106.4059
18.4104 | -106.4092
18.4141 | -106.4129
18.4177 | -106.4197
18.4214 | -106.4288
18.4251 | -106.4347
18.4288 | -106.4397
18.4324 | -106.4493
18.4361 | -106.4607
18.4398 | -106.4711
18.4435 | -106.4831
18.4471 | -106.4938
18.4508 | -106.5057
18.4545 | -106.5167
18.4581 | -106.528
18.4618 | -106.5412
18.4655 | -106.5548
18.4692 | -106.5687
18.4728 | -106.5855
18.4765 | -106.6014
18.4802 | -106.6152
18.4839 | -106.6314
18.4875 | -106.6473
18.4912 | -106.6633
18.4949 | -106.6776
18.4986 | -106.6919
18.5022 | -106.709
18.5059 | -106.7275
18.5096 | -106.7475
18.5132 | -106.7674
18.5169 | -106.7828
18.5206 | -106.7973
18.5243 | -106.8133
18.5279 | -106.8276
18.5316 | -106.8412
18.5353 | -106.8555
18.539 | -106.8683
18.5426 | -106.8813
18.5463 | -106.8932
18.55 | -106.9075
18.5537 | -106.9225
18.5573 | -106.9341
18.561 | -106.9428
18.5647 | -106.9546
18.5683 | -106.965
18.572 | -106.9722
18.5757 | -106.9764
18.5794 | -106.9814
18.583 | -106.9881
18.5867 | -106.991
18.5904 | -106.9891
18.5941 | -106.9861
18.5977 | -106.9821
18.6014 | -106.9766
18.6051 | -106.9708
18.6088 | -106.9667
18.6124 | -106.96
18.6161 | -106.9502
18.6198 | -106.9389
18.6234 | -106.927
18.6271 | -106.9144
18.6308 | -106.9006
18.6345 | -106.8848
18.6381 | -106.87
18.6418 | -106.8543
18.6455 | -106.8391
18.6492 | -106.8238
18.6528 | -106.8091
18.6565 | -106.7934
18.6602 | -106.7764
18.6639 | -106.758
18.6675 | -106.7402
18.6712 | -106.7271
18.6749 | -106.7141
18.6785 | -106.6987
18.6822 | -106.6824
18.6859 | -106.6677
18.6896 | -106.6563
18.6932 | -106.6441
18.6969 | -106.632
18.7006 | -106.6219
18.7043 | -106.6138
18.7079 | -106.6084
18.7116 | -106.6042
18.7153 | -106.5993
18.7189 | -106.5972
18.7226 | -106.5955
18.7263 | -106.5924
18.73 | -106.5889
18.7336 | -106.5875
18.7373 | -106.5894
18.741 | -106.5926
18.7447 | -106.5969
18.7483 | -106.6027
18.752 | -106.6094
18.7557 | -106.6176
18.7594 | -106.6287
18.763 | -106.64
18.7667 | -106.6489
18.7704 | -106.6586
18.774 | -106.6685
18.7777 | -106.6802
18.7814 | -106.6954
18.7851 | -106.7083
18.7887 | -106.7197
18.7924 | -106.7312
18.7961 | -106.7443
18.7998 | -106.7574
18.8034 | -106.7712
18.8071 | -106.786
18.8108 | -106.7978
18.8145 | -106.8097
18.8181 | -106.8233
18.8218 | -106.8374
18.8255 | -106.851
18.8291 | -106.8629
18.8328 | -106.8703
18.8365 | -106.879
18.8402 | -106.8903
18.8438 | -106.9012
18.8475 | -106.9114
18.8512 | -106.9185
18.8549 | -106.9291
18.8585 | -106.9422
18.8622 | -106.9538
18.8659 | -106.9641
18.8696 | -106.973
18.8732 | -106.9787
18.8769 | -106.9825
18.8806 | -106.9888
18.8842 | -106.9946
18.8879 | -106.9982
18.8916 | -106.9998
18.8953 | -107.0014
18.8989 | -107.0011
18.9026 | -107.0011
18.9063 | -106.9985
18.91 | -106.998
18.9136 | -106.9984
18.9173 | -106.9974
18.921 | -106.9955
18.9247 | -106.9941
18.9283 | -106.9938
18.932 | -106.9928
18.9357 | -106.9904
18.9393 | -106.9859
18.943 | -106.9807
18.9467 | -106.9786
18.9504 | -106.9768
18.954 | -106.975
18.9577 | -106.9705
18.9614 | -106.9695
18.9651 | -106.9661
18.9687 | -106.9612
18.9724 | -106.9569
18.9761 | -106.9527
18.9798 | -106.948
18.9834 | -106.9442
18.9871 | -106.9409
18.9908 | -106.9383
18.9944 | -106.9372
18.9981 | -106.9374
19.0018 | -106.9367
19.0055 | -106.9328
19.0091 | -106.9302
19.0128 | -106.9282
19.0165 | -106.9258
19.0202 | -106.9249
19.0238 | -106.9243
19.0275 | -106.9233
19.0312 | -106.9227
19.0348 | -106.9235
19.0385 | -106.9242
19.0422 | -106.9266
19.0459 | -106.9287
19.0495 | -106.9311
19.0532 | -106.9359
19.0569 | -106.9421
19.0606 | -106.9493
19.0642 | -106.9562
19.0679 | -106.9614
19.0716 | -106.965
19.0753 | -106.9737
19.0789 | -106.9785
19.0826 | -106.9863
19.0863 | -106.9929
19.0899 | -106.9987
19.0936 | -107.0051
19.0973 | -107.0114
19.101 | -107.0189
19.1046 | -107.0276
19.1083 | -107.0357
19.112 | -107.0412
19.1157 | -107.0463
19.1193 | -107.0511
19.123 | -107.0558
19.1267 | -107.0579
19.1304 | -107.0592
19.134 | -107.0608
19.1377 | -107.0619
19.1414 | -107.0646
19.145 | -107.0689
19.1487 | -107.0699
19.1524 | -107.0686
19.1561 | -107.0671
19.1597 | -107.0666
19.1634 | -107.0662
19.1671 | -107.0628
19.1708 | -107.0581
19.1744 | -107.0521
19.1781 | -107.0479
19.1818 | -107.0423
19.1855 | -107.0356
19.1891 | -107.0279
19.1928 | -107.0197
19.1965 | -107.0126
19.2001 | -107.0047
19.2038 | -106.9944
19.2075 | -106.9837
19.2112 | -106.9736
19.2148 | -106.9645
19.2185 | -106.9528
19.2222 | -106.941
19.2259 | -106.9293
19.2295 | -106.9192
19.2332 | -106.9097
19.2369 | -106.9018
19.2406 | -106.8936
19.2442 | -106.8836
19.2479 | -106.874
19.2516 | -106.8657
19.2552 | -106.8586
19.2589 | -106.8523
19.2626 | -106.8427
19.2663 | -106.834
19.2699 | -106.8265
19.2736 | -106.8188
19.2773 | -106.8088
19.281 | -106.8008
19.2846 | -106.7938
19.2883 | -106.7863
19.292 | -106.7797
19.2957 | -106.7743
19.2993 | -106.7693
19.303 | -106.7643
19.3067 | -106.7591
19.3103 | -106.7541
19.314 | -106.7481
19.3177 | -106.7428
19.3214 | -106.7369
19.325 | -106.7309
19.3287 | -106.7257
19.3324 | -106.7199
19.3361 | -106.7119
19.3397 | -106.7058
19.3434 | -106.7012
19.3471 | -106.6935
19.3508 | -106.6859
19.3544 | -106.6795
19.3581 | -106.6718
19.3618 | -106.6643
19.3654 | -106.658
19.3691 | -106.6516
19.3728 | -106.6458
19.3765 | -106.6396
19.3801 | -106.634
19.3838 | -106.6281
19.3875 | -106.6225
19.3912 | -106.6152
19.3948 | -106.6086
19.3985 | -106.605
19.4022 | -106.6025
19.4059 | -106.597
19.4095 | -106.5919
19.4132 | -106.5877
19.4169 | -106.5817
19.4205 | -106.578
19.4242 | -106.5743
19.4279 | -106.5697
19.4316 | -106.5673
19.4352 | -106.5638
19.4389 | -106.5608
19.4426 | -106.5595
19.4463 | -106.558
19.4499 | -106.5545
19.4536 | -106.5487
19.4573 | -106.5461
19.4609 | -106.5437
19.4646 | -106.5417
19.4683 | -106.5407
19.472 | -106.5408
19.4756 | -106.539
19.4793 | -106.5382
19.483 | -106.5355
19.4867 | -106.5319
19.4903 | -106.5304
19.494 | -106.5315
19.4977 | -106.5346
19.5014 | -106.536
19.505 | -106.5397
19.5087 | -106.5445
19.5124 | -106.5479
19.516 | -106.5524
19.5197 | -106.5575
19.5234 | -106.5627
19.5271 | -106.5704
19.5307 | -106.5787
19.5344 | -106.5878
19.5381 | -106.5994
19.5418 | -106.6078
19.5454 | -106.6143
19.5491 | -106.6231
19.5528 | -106.6342
19.5565 | -106.6431
19.5601 | -106.6491
19.5638 | -106.6542
19.5675 | -106.6587
19.5711 | -106.6622
19.5748 | -106.6659
19.5785 | -106.6705
19.5822 | -106.6738
19.5858 | -106.6761
19.5895 | -106.6785
19.5932 | -106.6804
19.5969 | -106.6824
19.6005 | -106.6839
19.6042 | -106.6762
19.6079 | -106.6726
19.6116 | -106.6677
19.6152 | -106.6657
19.6189 | -106.6623
19.6226 | -106.6597
19.6262 | -106.6578
19.6299 | -106.6536
19.6336 | -106.6491
19.6373 | -106.6443
19.6409 | -106.6387
19.6446 | -106.631
19.6483 | -106.6236
19.652 | -106.616
19.6556 | -106.608
19.6593 | -106.6
19.663 | -106.5911
19.6667 | -106.5838
19.6703 | -106.5777
19.674 | -106.5701
19.6777 | -106.5617
19.6813 | -106.5545
19.685 | -106.5496
19.6887 | -106.544
19.6924 | -106.5393
19.696 | -106.5345
19.6997 | -106.5288
19.7034 | -106.5252
19.7071 | -106.5237
19.7107 | -106.5208
19.7144 | -106.5176
19.7181 | -106.5147
19.7218 | -106.5139
19.7254 | -106.5155
19.7291 | -106.519
19.7328 | -106.5247
19.7364 | -106.5291
19.7401 | -106.5336
19.7438 | -106.5401
19.7475 | -106.5468
19.7511 | -106.5517
19.7548 | -106.5581
19.7585 | -106.5674
19.7622 | -106.58
19.7658 | -106.5923
19.7695 | -106.6038
19.7732 | -106.6144
19.7768 | -106.6236
19.7805 | -106.6343
19.7842 | -106.644
19.7879 | -106.6521
19.7915 | -106.6619
19.7952 | -106.6724
19.7989 | -106.6844
19.8026 | -106.6937
19.8062 | -106.6983
19.8099 | -106.7019
19.8136 | -106.7065
19.8173 | -106.7107
19.8209 | -106.7147
19.8246 | -106.7197
19.8283 | -106.7239
19.8319 | -106.7282
19.8356 | -106.7309
19.8393 | -106.7339
19.843 | -106.7368
19.8466 | -106.7368
19.8503 | -106.7375
19.854 | -106.7384
19.8577 | -106.7365
19.8613 | -106.7348
19.865 | -106.7304
19.8687 | -106.7252
19.8724 | -106.722
19.876 | -106.7181
19.8797 | -106.7144
19.8834 | -106.7102
19.887 | -106.7061
19.8907 | -106.7045
19.8944 | -106.7015
19.8981 | -106.695
19.9017 | -106.6887
19.9054 | -106.6808
19.9091 | -106.6735
19.9128 | -106.6693
19.9164 | -106.6645
19.9201 | -106.6607
19.9238 | -106.6573
19.9275 | -106.6561
19.9311 | -106.6546
19.9348 | -106.6535
19.9385 | -106.6537
19.9421 | -106.6528
19.9458 | -106.6513
19.9495 | -106.6499
19.9532 | -106.6486
19.9568 | -106.6485
19.9605 | -106.6454
19.9642 | -106.6436
19.9679 | -106.6432
19.9715 | -106.6436
19.9752 | -106.6447
19.9789 | -106.6453
19.9826 | -106.6465
19.9862 | -106.6473
19.9899 | -106.645
19.9936 | -106.6435
19.9972 | -106.6451
20.0009 | -106.6457
20.0046 | -106.6454
20.0083 | -106.6447
20.0119 | -106.6455
20.0156 | -106.6468
20.0193 | -106.6465
20.023 | -106.6449
20.0266 | -106.6434
20.0303 | -106.6425
20.034 | -106.6423
20.0377 | -106.6448
20.0413 | -106.6478
20.045 | -106.6479
20.0487 | -106.6471
20.0523 | -106.6472
20.056 | -106.6466
20.0597 | -106.6429
20.0634 | -106.6433
20.067 | -106.6447
20.0707 | -106.6458
20.0744 | -106.6482
20.0781 | -106.6503
20.0817 | -106.6517
20.0854 | -106.6517
20.0891 | -106.6538
20.0927 | -106.6571
20.0964 | -106.6576
20.1001 | -106.6584
20.1038 | -106.6608
20.1074 | -106.6596
20.1111 | -106.6573
20.1148 | -106.6558
20.1185 | -106.6561
20.1221 | -106.6533
20.1258 | -106.6503
20.1295 | -106.6495
20.1332 | -106.6471
20.1368 | -106.6406
20.1405 | -106.6347
20.1442 | -106.6317
20.1478 | -106.627
20.1515 | -106.6243
20.1552 | -106.6223
20.1589 | -106.6178
20.1625 | -106.6119
20.1662 | -106.6066
20.1699 | -106.603
20.1736 | -106.5991
20.1772 | -106.595
20.1809 | -106.5898
20.1846 | -106.5847
20.1883 | -106.5816
20.1919 | -106.5777
20.1956 | -106.5731
20.1993 | -106.5696
20.2029 | -106.5642
20.2066 | -106.5568
20.2103 | -106.553
20.214 | -106.55
20.2176 | -106.5476
20.2213 | -106.5459
20.225 | -106.5433
20.2287 | -106.5416
20.2323 | -106.542
20.236 | -106.5416
20.2397 | -106.5422
20.2434 | -106.5425
20.247 | -106.5434
20.2507 | -106.5425
20.2544 | -106.5417
20.258 | -106.5408
20.2617 | -106.5373
20.2654 | -106.5345
20.2691 | -106.5302
20.2727 | -106.5286
20.2764 | -106.5282
20.2801 | -106.525
20.2838 | -106.5208
20.2874 | -106.5174
20.2911 | -106.5118
20.2948 | -106.5068
20.2985 | -106.5016
20.3021 | -106.497
20.3058 | -106.4924
20.3095 | -106.4908
20.3131 | -106.4885
20.3168 | -106.4859
20.3205 | -106.4828
20.3242 | -106.4777
20.3278 | -106.4725
20.3315 | -106.4668
20.3352 | -106.4634
20.3389 | -106.46
20.3425 | -106.4588
20.3462 | -106.4608
20.3499 | -106.4583
20.3536 | -106.4553
20.3572 | -106.454
20.3609 | -106.4526
20.3646 | -106.4524
20.3682 | -106.4529
20.3719 | -106.4565
20.3756 | -106.46
20.3793 | -106.4633
20.3829 | -106.4665
20.3866 | -106.4704
20.3903 | -106.4735
20.394 | -106.4775
20.3976 | -106.4826
20.4013 | -106.4861
20.405 | -106.4879
20.4087 | -106.4943
20.4123 | -106.5007
20.416 | -106.5065
20.4197 | -106.5128
20.4233 | -106.52
20.427 | -106.5255
20.4307 | -106.5321
20.4344 | -106.5377
20.438 | -106.5464
20.4417 | -106.5541
20.4454 | -106.5583
20.4491 | -106.5631
20.4527 | -106.5694
20.4564 | -106.5752
20.4601 | -106.5802
20.4637 | -106.5847
20.4674 | -106.5874
20.4711 | -106.5926
20.4748 | -106.5991
20.4784 | -106.6023
20.4821 | -106.606
20.4858 | -106.609
20.4895 | -106.6071
20.4931 | -106.6074
20.4968 | -106.6091
20.5005 | -106.6055
20.5042 | -106.6008
20.5078 | -106.5969
20.5115 | -106.5937
20.5152 | -106.5885
20.5188 | -106.5823
20.5225 | -106.576
20.5262 | -106.5704
20.5299 | -106.5649
20.5335 | -106.5573
20.5372 | -106.5489
20.5409 | -106.5422
20.5446 | -106.5348
20.5482 | -106.5268
20.5519 | -106.5191
20.5556 | -106.5118
20.5593 | -106.5046
20.5629 | -106.4952
20.5666 | -106.4869
20.5703 | -106.4794
20.5739 | -106.471
20.5776 | -106.4605
20.5813 | -106.4517
20.585 | -106.444
20.5886 | -106.4365
20.5923 | -106.4292
20.596 | -106.4215
20.5997 | -106.4119
20.6033 | -106.4024
20.607 | -106.3932
20.6107 | -106.3868
20.6144 | -106.38
20.618 | -106.3724
20.6217 | -106.3669
20.6254 | -106.3624
20.629 | -106.358
20.6327 | -106.3529
20.6364 | -106.3468
20.6401 | -106.3418
20.6437 | -106.3379
20.6474 | -106.3328
20.6511 | -106.3275
20.6548 | -106.3198
20.6584 | -106.3137
20.6621 | -106.3085
20.6658 | -106.31
20.6695 | -106.307
20.6731 | -106.3025
20.6768 | -106.3001
20.6805 | -106.2971
20.6841 | -106.2942
20.6878 | -106.2937
20.6915 | -106.292
20.6952 | -106.291
20.6988 | -106.2906
20.7025 | -106.2894
20.7062 | -106.2875
20.7099 | -106.2841
20.7135 | -106.2811
20.7172 | -106.2764
20.7209 | -106.2738
20.7246 | -106.2725
20.7282 | -106.2698
20.7319 | -106.2657
20.7356 | -106.2631
20.7392 | -106.2618
20.7429 | -106.2588
20.7466 | -106.2533
20.7503 | -106.2479
20.7539 | -106.2448
20.7576 | -106.2409
20.7613 | -106.2367
20.765 | -106.2316
20.7686 | -106.2262
20.7723 | -106.2222
20.776 | -106.2182
20.7797 | -106.2137
20.7833 | -106.2107
20.787 | -106.2081
20.7907 | -106.2049
20.7943 | -106.2022
20.798 | -106.1999
20.8017 | -106.1958
20.8054 | -106.1903
20.809 | -106.1871
20.8127 | -106.1852
20.8164 | -106.183
20.8201 | -106.1787
20.8237 | -106.1741
20.8274 | -106.1726
20.8311 | -106.1727
20.8347 | -106.1726
20.8384 | -106.1724
20.8421 | -106.1738
20.8458 | -106.1745
20.8494 | -106.1727
20.8531 | -106.1729
20.8568 | -106.1738
20.8605 | -106.1757
20.8641 | -106.1792
20.8678 | -106.1835
20.8715 | -106.1886
20.8752 | -106.1937
20.8788 | -106.1975
20.8825 | -106.2029
20.8862 | -106.2089
20.8898 | -106.2164
20.8935 | -106.224
20.8972 | -106.2312
20.9009 | -106.2397
20.9045 | -106.2445
20.9082 | -106.2507
20.9119 | -106.2581
20.9156 | -106.2643
20.9192 | -106.2691
20.9229 | -106.2744
20.9266 | -106.2802
20.9303 | -106.2864
20.9339 | -106.2914
20.9376 | -106.2948
20.9413 | -106.2995
20.9449 | -106.3036
20.9486 | -106.3077
20.9523 | -106.3087
20.956 | -106.3098
20.9596 | -106.3069
20.9633 | -106.3039
20.967 | -106.302
20.9707 | -106.299
20.9743 | -106.293
20.978 | -106.288
20.9817 | -106.283
20.9854 | -106.2741
20.989 | -106.2647
20.9927 | -106.2553
20.9964 | -106.2463
21 | -106.2363
21.0037 | -106.2258
21.0074 | -106.2151
21.0111 | -106.2058
21.0147 | -106.1987
21.0184 | -106.1885
21.0221 | -106.1733
21.0258 | -106.1576
21.0294 | -106.1443
21.0331 | -106.132
21.0368 | -106.121
21.0405 | -106.1102
21.0441 | -106.1026
21.0478 | -106.0948
21.0515 | -106.0842
21.0551 | -106.0739
21.0588 | -106.0646
21.0625 | -106.0569
21.0662 | -106.0483
21.0698 | -106.0401
21.0735 | -106.0335
21.0772 | -106.0278
21.0809 | -106.0222
21.0845 | -106.0178
21.0882 | -106.0138
21.0919 | -106.0118
21.0956 | -106.0097
21.0992 | -106.0082
21.1029 | -106.0088
21.1066 | -106.0096
21.1102 | -106.0086
21.1139 | -106.0079
21.1176 | -106.0104
21.1213 | -106.0148
21.1249 | -106.02
21.1286 | -106.0258
21.1323 | -106.0286
21.136 | -106.0335
21.1396 | -106.0396
21.1433 | -106.0482
21.147 | -106.0578
21.1506 | -106.0663
21.1543 | -106.0753
21.158 | -106.0847
21.1617 | -106.0929
21.1653 | -106.1005
21.169 | -106.1099
21.1727 | -106.1205
21.1764 | -106.1322
21.18 | -106.1422
21.1837 | -106.1526
21.1874 | -106.1641
21.1911 | -106.1767
21.1947 | -106.1871
21.1984 | -106.2
21.2021 | -106.2108
21.2057 | -106.2222
21.2094 | -106.2339
21.2131 | -106.2466
21.2168 | -106.2599
21.2204 | -106.2721
21.2241 | -106.2817
21.2278 | -106.2884
21.2315 | -106.2965
21.2351 | -106.307
21.2388 | -106.3179
21.2425 | -106.3266
21.2462 | -106.3336
21.2498 | -106.3424
21.2535 | -106.3518
21.2572 | -106.3584
21.2608 | -106.3634
21.2645 | -106.368
21.2682 | -106.371
21.2719 | -106.3763
21.2755 | -106.382
21.2792 | -106.3869
21.2829 | -106.39
21.2866 | -106.3936
21.2902 | -106.3968
21.2939 | -106.4005
21.2976 | -106.4035
21.3013 | -106.4061
21.3049 | -106.4071
21.3086 | -106.4082
21.3123 | -106.4085
21.3159 | -106.4087
21.3196 | -106.4095
21.3233 | -106.4117
21.327 | -106.4143
21.3306 | -106.4146
21.3343 | -106.413
21.338 | -106.4107
21.3417 | -106.4116
21.3453 | -106.4113
21.349 | -106.4097
21.3527 | -106.4103
21.3564 | -106.4098
21.36 | -106.407
21.3637 | -106.4052
21.3674 | -106.4045
21.371 | -106.4015
21.3747 | -106.3997
21.3784 | -106.4003
21.3821 | -106.4009
21.3857 | -106.3991
21.3894 | -106.3972
21.3931 | -106.3954
21.3968 | -106.3931
21.4004 | -106.3913
21.4041 | -106.3892
21.4078 | -106.3881
21.4115 | -106.3884
21.4151 | -106.3882
21.4188 | -106.3863
21.4225 | -106.3809
21.4261 | -106.3733
21.4298 | -106.3677
21.4335 | -106.3645
21.4372 | -106.3595
21.4408 | -106.3522
21.4445 | -106.347
21.4482 | -106.3436
21.4519 | -106.34
21.4555 | -106.3349
21.4592 | -106.3275
21.4629 | -106.3225
21.4666 | -106.3185
21.4702 | -106.3125
21.4739 | -106.3052
21.4776 | -106.2973
21.4812 | -106.2914
21.4849 | -106.2854
21.4886 | -106.2767
21.4923 | -106.2689
21.4959 | -106.2628
21.4996 | -106.2572
21.5033 | -106.2496
21.507 | -106.2417
21.5106 | -106.2355
21.5143 | -106.226
21.518 | -106.216
21.5216 | -106.207
21.5253 | -106.1998
21.529 | -106.1941
21.5327 | -106.1914
21.5363 | -106.1866
21.54 | -106.1799
21.5437 | -106.1729
21.5474 | -106.1681
21.551 | -106.1636
21.5547 | -106.1574
21.5584 | -106.1504
21.5621 | -106.1462
21.5657 | -106.144
21.5694 | -106.1428
21.5731 | -106.1416
21.5767 | -106.1387
21.5804 | -106.1368
21.5841 | -106.1342
21.5878 | -106.132
21.5914 | -106.1316
21.5951 | -106.1278
21.5988 | -106.1221
21.6025 | -106.1182
21.6061 | -106.1141
21.6098 | -106.1124
21.6135 | -106.1081
21.6172 | -106.1056
21.6208 | -106.1026
21.6245 | -106.0974
21.6282 | -106.0939
21.6318 | -106.092
21.6355 | -106.0879
21.6392 | -106.0823
21.6429 | -106.0771
21.6465 | -106.0706
21.6502 | -106.0643
21.6539 | -106.0587
21.6576 | -106.0517
21.6612 | -106.0456
21.6649 | -106.0406
21.6686 | -106.0337
21.6723 | -106.0269
21.6759 | -106.0191
21.6796 | -106.0104
21.6833 | -106.0003
21.6869 | -105.9925
21.6906 | -105.9858
21.6943 | -105.9765
21.698 | -105.9664
21.7016 | -105.9566
21.7053 | -105.9494
21.709 | -105.942
21.7127 | -105.9331
21.7163 | -105.9254
21.72 | -105.9187
21.7237 | -105.9119
21.7274 | -105.9084
21.731 | -105.9036
21.7347 | -105.9006
21.7384 | -105.8959
21.742 | -105.8914
21.7457 | -105.8873
21.7494 | -105.8861
21.7531 | -105.8829
21.7567 | -105.8806
21.7604 | -105.8793
21.7641 | -105.8778
21.7678 | -105.8766
21.7714 | -105.8765
21.7751 | -105.8779
21.7788 | -105.8811
21.7825 | -105.8839
21.7861 | -105.8875
21.7898 | -105.8895
21.7935 | -105.8928
21.7971 | -105.8985
21.8008 | -105.9048
21.8045 | -105.9098
21.8082 | -105.9139
21.8118 | -105.9208
21.8155 | -105.9282
21.8192 | -105.9354
21.8229 | -105.9453
21.8265 | -105.9544
21.8302 | -105.9628
21.8339 | -105.9694
21.8376 | -105.9754
21.8412 | -105.9817
21.8449 | -105.9896
21.8486 | -105.9977
21.8522 | -106.0039
21.8559 | -106.0095
21.8596 | -106.0151
21.8633 | -106.0215
21.8669 | -106.03
21.8706 | -106.0366
21.8743 | -106.0407
21.878 | -106.0424
21.8816 | -106.0435
21.8853 | -106.0443
21.889 | -106.0445
21.8926 | -106.0446
21.8963 | -106.046
21.9 | -106.0454
21.9037 | -106.0475
21.9073 | -106.048
21.911 | -106.0457
21.9147 | -106.0433
21.9184 | -106.0392
21.922 | -106.0377
21.9257 | -106.0325
21.9294 | -106.0275
21.9331 | -106.025
21.9367 | -106.021
21.9404 | -106.0135
21.9441 | -106.0074
21.9477 | -106.0033
21.9514 | -105.9988
21.9551 | -105.994
21.9588 | -105.9871
21.9624 | -105.9814
21.9661 | -105.9765
21.9698 | -105.9707
21.9735 | -105.965
21.9771 | -105.9588
21.9808 | -105.9515
21.9845 | -105.9468
21.9882 | -105.9407
21.9918 | -105.9357
21.9955 | -105.9287
21.9992 | -105.9223
22.0028 | -105.9187
22.0065 | -105.9162
22.0102 | -105.9129
22.0139 | -105.9083
22.0175 | -105.9048
22.0212 | -105.9024
22.0249 | -105.8991
22.0286 | -105.8946
22.0322 | -105.8918
22.0359 | -105.8896
22.0396 | -105.8883
22.0433 | -105.8878
22.0469 | -105.8866
22.0506 | -105.8858
22.0543 | -105.8859
22.0579 | -105.8879
22.0616 | -105.8881
22.0653 | -105.8883
22.069 | -105.8892
22.0726 | -105.8905
22.0763 | -105.8927
22.08 | -105.897
22.0837 | -105.9
22.0873 | -105.9011
22.091 | -105.9034
22.0947 | -105.9079
22.0984 | -105.913
22.102 | -105.918
22.1057 | -105.9221
22.1094 | -105.9279
22.113 | -105.9328
22.1167 | -105.9379
22.1204 | -105.9451
22.1241 | -105.9502
22.1277 | -105.9563
22.1314 | -105.962
22.1351 | -105.967
22.1388 | -105.9729
22.1424 | -105.9797
22.1461 | -105.9864
22.1498 | -105.9935
22.1535 | -106.0013
22.1571 | -106.0065
22.1608 | -106.013
22.1645 | -106.0215
22.1681 | -106.0288
22.1718 | -106.0366
22.1755 | -106.044
22.1792 | -106.0498
22.1828 | -106.0552
22.1865 | -106.0632
22.1902 | -106.0734
22.1939 | -106.0833
22.1975 | -106.0933
22.2012 | -106.101
22.2049 | -106.109
22.2085 | -106.117
22.2122 | -106.1258
22.2159 | -106.134
22.2196 | -106.1435
22.2232 | -106.1538
22.2269 | -106.1638
22.2306 | -106.1732
22.2343 | -106.1812
22.2379 | -106.1898
22.2416 | -106.1986
22.2453 | -106.2074
22.249 | -106.2153
22.2526 | -106.2235
22.2563 | -106.231
22.26 | -106.2382
22.2636 | -106.2462
22.2673 | -106.2538
22.271 | -106.2603
22.2747 | -106.2681
22.2783 | -106.2753
22.282 | -106.2833
22.2857 | -106.2915
22.2894 | -106.2974
22.293 | -106.3046
22.2967 | -106.3128
22.3004 | -106.3201
22.3041 | -106.3272
22.3077 | -106.3329
22.3114 | -106.3382
22.3151 | -106.3433
22.3187 | -106.3477
22.3224 | -106.3494
22.3261 | -106.3531
22.3298 | -106.356
22.3334 | -106.3579
22.3371 | -106.3602
22.3408 | -106.3629
22.3445 | -106.3676
22.3481 | -106.3725
22.3518 | -106.3758
22.3555 | -106.3777
22.3592 | -106.3784
22.3628 | -106.3796
22.3665 | -106.3793
22.3702 | -106.379
22.3738 | -106.3759
22.3775 | -106.3741
22.3812 | -106.3716
22.3849 | -106.3674
22.3885 | -106.3621
22.3922 | -106.3539
22.3959 | -106.342
22.3996 | -106.3296
22.4032 | -106.318
22.4069 | -106.3037
22.4106 | -106.2879
22.4143 | -106.2714
22.4179 | -106.2567
22.4216 | -106.2362
22.4253 | -106.2151
22.4289 | -106.1938
22.4326 | -106.1725
22.4363 | -106.1471
22.44 | -106.1224
22.4436 | -106.0949
22.4473 | -106.0667
22.451 | -106.0377
22.4547 | -106.0048
22.4583 | -105.9715
22.462 | -105.9394
22.4657 | -105.9068
22.4694 | -105.8717
22.473 | -105.8363
22.4767 | -105.7988
22.4804 | -105.7605
22.484 | -105.722
22.4877 | -105.6833
22.4914 | -105.644
22.4951 | -105.6036
22.4987 | -105.5647
22.5024 | -105.5255
22.5061 | -105.4863
22.5098 | -105.4485
22.5134 | -105.4115
22.5171 | -105.3774
22.5208 | -105.3427
22.5245 | -105.3109
22.5281 | -105.2817
22.5318 | -105.254
22.5355 | -105.228
22.5391 | -105.2037
22.5428 | -105.1828
22.5465 | -105.1655
22.5502 | -105.1511
22.5538 | -105.1384
22.5575 | -105.1281
22.5612 | -105.1214
22.5649 | -105.1177
22.5685 | -105.1181
22.5722 | -105.1205
22.5759 | -105.1269
22.5795 | -105.137
22.5832 | -105.1486
22.5869 | -105.1632
22.5906 | -105.1807
22.5942 | -105.202
22.5979 | -105.2262
22.6016 | -105.2523
22.6053 | -105.2801
22.6089 | -105.3106
22.6126 | -105.3415
22.6163 | -105.3745
22.62 | -105.4081
22.6236 | -105.4436
22.6273 | -105.4821
22.631 | -105.5234
22.6346 | -105.5648
22.6383 | -105.6069
22.642 | -105.6501
22.6457 | -105.6924
22.6493 | -105.7346
22.653 | -105.777
22.6567 | -105.8187
22.6604 | -105.8611
22.664 | -105.9016
22.6677 | -105.9387
22.6714 | -105.9763
22.6751 | -106.0157
22.6787 | -106.0534
22.6824 | -106.0903
22.6861 | -106.1269
22.6897 | -106.161
22.6934 | -106.1944
22.6971 | -106.2262
22.7008 | -106.2582
22.7044 | -106.2884
22.7081 | -106.3178
22.7118 | -106.3468
22.7155 | -106.3763
22.7191 | -106.403
22.7228 | -106.4267
22.7265 | -106.4489
22.7302 | -106.4684
22.7338 | -106.4854
22.7375 | -106.4998
22.7412 | -106.5166
22.7448 | -106.5285
22.7485 | -106.5373
22.7522 | -106.5418
22.7559 | -106.5497
22.7595 | -106.5549
22.7632 | -106.5611
22.7669 | -106.566
22.7706 | -106.5697
22.7742 | -106.5737
22.7779 | -106.5756
22.7816 | -106.5759
22.7853 | -106.5732
22.7889 | -106.5717
22.7926 | -106.5696
22.7963 | -106.5674
22.7999 | -106.5656
22.8036 | -106.5623
22.8073 | -106.5603
22.811 | -106.5554
22.8146 | -106.552
22.8183 | -106.5499
22.822 | -106.5471
22.8257 | -106.5424
22.8293 | -106.538
22.833 | -106.5335
22.8367 | -106.5275
22.8404 | -106.5207
22.844 | -106.5149
22.8477 | -106.5083
22.8514 | -106.5004
22.855 | -106.4944
22.8587 | -106.4875
22.8624 | -106.4789
22.8661 | -106.4721
22.8697 | -106.4663
22.8734 | -106.4593
22.8771 | -106.4524
22.8808 | -106.4471
22.8844 | -106.442
22.8881 | -106.4374
22.8918 | -106.4322
22.8955 | -106.4261
22.8991 | -106.4201
22.9028 | -106.4138
22.9065 | -106.4081
22.9101 | -106.403
22.9138 | -106.3988
22.9175 | -106.3936
22.9212 | -106.3882
22.9248 | -106.384
22.9285 | -106.3807
22.9322 | -106.378
22.9359 | -106.3768
22.9395 | -106.3757
22.9432 | -106.3755
22.9469 | -106.3739
22.9505 | -106.371
22.9542 | -106.3698
22.9579 | -106.3698
22.9616 | -106.3697
22.9652 | -106.3697
22.9689 | -106.3693
22.9726 | -106.3696
22.9763 | -106.3698
22.9799 | -106.3714
22.9836 | -106.3709
22.9873 | -106.371
22.991 | -106.3728
22.9946 | -106.3752
22.9983 | -106.3771
23.002 | -106.3808
23.0056 | -106.3821
23.0093 | -106.3847
23.013 | -106.3889
23.0167 | -106.3925
23.0203 | -106.3959
23.024 | -106.4
23.0277 | -106.4044
23.0314 | -106.4085
23.035 | -106.4158
23.0387 | -106.4208
23.0424 | -106.4239
23.0461 | -106.4272
23.0497 | -106.4305
23.0534 | -106.4355
23.0571 | -106.4404
23.0607 | -106.4443
23.0644 | -106.449
23.0681 | -106.454
23.0718 | -106.4603
23.0754 | -106.4662
23.0791 | -106.4713
23.0828 | -106.4763
23.0865 | -106.4793
23.0901 | -106.481
23.0938 | -106.4822
23.0975 | -106.4847
23.1012 | -106.4867
23.1048 | -106.4895
23.1085 | -106.4916
23.1122 | -106.4939
23.1158 | -106.4964
23.1195 | -106.4991
23.1232 | -106.5023
23.1269 | -106.5049
23.1305 | -106.5061
23.1342 | -106.5058
23.1379 | -106.5077
23.1416 | -106.5071
23.1452 | -106.5069
23.1489 | -106.5063
23.1526 | -106.5058
23.1563 | -106.5035
23.1599 | -106.5019
23.1636 | -106.5012
23.1673 | -106.502
23.1709 | -106.5033
23.1746 | -106.5006
23.1783 | -106.4994
23.182 | -106.498
23.1856 | -106.4964
23.1893 | -106.4952
23.193 | -106.4916
23.1967 | -106.4873
23.2003 | -106.485
23.204 | -106.4841
23.2077 | -106.48
23.2113 | -106.4739
23.215 | -106.4684
23.2187 | -106.4625
23.2224 | -106.4567
23.226 | -106.4526
23.2297 | -106.4506
23.2334 | -106.4476
23.2371 | -106.4426
23.2407 | -106.4374
23.2444 | -106.4326
23.2481 | -106.4287
23.2518 | -106.4247
23.2554 | -106.4207
23.2591 | -106.4172
23.2628 | -106.4123
23.2664 | -106.4076
23.2701 | -106.4039
23.2738 | -106.3995
23.2775 | -106.3946
23.2811 | -106.3903
23.2848 | -106.3864
23.2885 | -106.3829
23.2922 | -106.3811
23.2958 | -106.3788
23.2995 | -106.3762
23.3032 | -106.374
23.3069 | -106.3708
23.3105 | -106.3698
23.3142 | -106.3701
23.3179 | -106.3699
23.3215 | -106.3694
23.3252 | -106.3684
23.3289 | -106.3673
23.3326 | -106.3674
23.3362 | -106.3691
23.3399 | -106.3691
23.3436 | -106.3683
23.3473 | -106.3715
23.3509 | -106.3736
23.3546 | -106.3753
23.3583 | -106.3781
23.362 | -106.3809
23.3656 | -106.3821
23.3693 | -106.3851
23.373 | -106.3903
23.3766 | -106.3963
23.3803 | -106.4025
23.384 | -106.4078
23.3877 | -106.4129
23.3913 | -106.417
23.395 | -106.4218
23.3987 | -106.4274
23.4024 | -106.4326
23.406 | -106.4404
23.4097 | -106.4465
23.4134 | -106.4514
23.4171 | -106.4575
23.4207 | -106.4631
23.4244 | -106.469
23.4281 | -106.4733
23.4317 | -106.4789
23.4354 | -106.4843
23.4391 | -106.4903
23.4428 | -106.4971
23.4464 | -106.5042
23.4501 | -106.5117
23.4538 | -106.5185
23.4575 | -106.5244
23.4611 | -106.5327
23.4648 | -106.5412
23.4685 | -106.5488
23.4722 | -106.5579
23.4758 | -106.5649
23.4795 | -106.572
23.4832 | -106.5798
23.4868 | -106.5897
23.4905 | -106.5987
23.4942 | -106.6073
23.4979 | -106.6135
23.5015 | -106.6195
23.5052 | -106.6267
23.5089 | -106.6346
23.5126 | -106.6429
23.5162 | -106.6522
23.5199 | -106.6616
23.5236 | -106.6713
23.5273 | -106.6809
23.5309 | -106.6887
23.5346 | -106.6955
23.5383 | -106.7039
23.5419 | -106.7128
23.5456 | -106.7232
23.5493 | -106.7335
23.553 | -106.7435
23.5566 | -106.7501
23.5603 | -106.7554
23.564 | -106.7612
23.5677 | -106.7676
23.5713 | -106.7752
23.575 | -106.783
23.5787 | -106.7873
23.5824 | -106.7915
23.586 | -106.7987
23.5897 | -106.8063
23.5934 | -106.8144
23.597 | -106.8199
23.6007 | -106.8245
23.6044 | -106.8303
23.6081 | -106.8376
23.6117 | -106.8419
23.6154 | -106.8462
23.6191 | -106.8501
23.6228 | -106.8535
23.6264 | -106.8588
23.6301 | -106.8631
23.6338 | -106.8679
23.6374 | -106.8731
23.6411 | -106.8779
23.6448 | -106.881
23.6485 | -106.8808
23.6521 | -106.8821
23.6558 | -106.8856
23.6595 | -106.8875
23.6632 | -106.8888
23.6668 | -106.8883
23.6705 | -106.8878
23.6742 | -106.889
23.6779 | -106.8895
23.6815 | -106.886
23.6852 | -106.8844
23.6889 | -106.8825
23.6925 | -106.8798
23.6962 | -106.8787
23.6999 | -106.8786
23.7036 | -106.8773
23.7072 | -106.8756
23.7109 | -106.8714
23.7146 | -106.8665
23.7183 | -106.8615
23.7219 | -106.8561
23.7256 | -106.8524
23.7293 | -106.8467
23.733 | -106.8395
23.7366 | -106.8343
23.7403 | -106.8283
23.744 | -106.8223
23.7476 | -106.8169
23.7513 | -106.8122
23.755 | -106.8102
23.7587 | -106.8073
23.7623 | -106.8039
23.766 | -106.7996
23.7697 | -106.7955
23.7734 | -106.792
23.777 | -106.7878
23.7807 | -106.7833
23.7844 | -106.7779
23.7881 | -106.7747
23.7917 | -106.7731
23.7954 | -106.7729
23.7991 | -106.7713
23.8027 | -106.7686
23.8064 | -106.7647
23.8101 | -106.7631
23.8138 | -106.7604
23.8174 | -106.7605
23.8211 | -106.7592
23.8248 | -106.757
23.8285 | -106.7558
23.8321 | -106.7548
23.8358 | -106.7551
23.8395 | -106.7567
23.8432 | -106.7551
23.8468 | -106.7554
23.8505 | -106.7547
23.8542 | -106.7535
23.8578 | -106.7512
23.8615 | -106.7491
23.8652 | -106.7462
23.8689 | -106.7438
23.8725 | -106.7438
23.8762 | -106.7436
23.8799 | -106.7444
23.8836 | -106.7413
23.8872 | -106.739
23.8909 | -106.7371
23.8946 | -106.737
23.8983 | -106.7366
23.9019 | -106.7365
23.9056 | -106.7372
23.9093 | -106.7337
23.9129 | -106.7304
23.9166 | -106.7312
23.9203 | -106.7297
23.924 | -106.7257
23.9276 | -106.7232
23.9313 | -106.719
23.935 | -106.7174
23.9387 | -106.716
23.9423 | -106.7142
23.946 | -106.7112
23.9497 | -106.7079
23.9534 | -106.7031
23.957 | -106.7007
23.9607 | -106.6982
23.9644 | -106.6938
23.968 | -106.6897
23.9717 | -106.6863
23.9754 | -106.6826
23.9791 | -106.6795
23.9827 | -106.6767
23.9864 | -106.673
23.9901 | -106.6693
23.9938 | -106.6648
23.9974 | -106.6605
24.0011 | -106.656
24.0048 | -106.6517
24.0084 | -106.6493
24.0121 | -106.6472
24.0158 | -106.6459
24.0195 | -106.6427
24.0231 | -106.6408
24.0268 | -106.6401
24.0305 | -106.639
24.0342 | -106.637
24.0378 | -106.6356
24.0415 | -106.6344
24.0452 | -106.6331
24.0489 | -106.6339
24.0525 | -106.6351
24.0562 | -106.6355
24.0599 | -106.637
24.0635 | -106.639
24.0672 | -106.6409
24.0709 | -106.6426
24.0746 | -106.6419
24.0782 | -106.6426
24.0819 | -106.6413
24.0856 | -106.6404
24.0893 | -106.6408
24.0929 | -106.6403
24.0966 | -106.6407
24.1003 | -106.6422
24.104 | -106.6435
24.1076 | -106.6438
24.1113 | -106.6432
24.115 | -106.6428
24.1186 | -106.6439
24.1223 | -106.6454
24.126 | -106.646
24.1297 | -106.6456
24.1333 | -106.6453
24.137 | -106.6459
24.1407 | -106.647
24.1444 | -106.6491
24.148 | -106.6519
24.1517 | -106.6551
24.1554 | -106.6555
24.1591 | -106.6558
24.1627 | -106.6574
24.1664 | -106.6589
24.1701 | -106.6609
24.1737 | -106.6628
24.1774 | -106.6644
24.1811 | -106.6676
24.1848 | -106.6709
24.1884 | -106.6726
24.1921 | -106.6749
24.1958 | -106.678
24.1995 | -106.6815
24.2031 | -106.6841
24.2068 | -106.688
24.2105 | -106.6929
24.2142 | -106.6957
24.2178 | -106.6983
24.2215 | -106.7024
24.2252 | -106.7084
24.2288 | -106.7123
24.2325 | -106.7167
24.2362 | -106.7201
24.2399 | -106.7218
24.2435 | -106.7248
24.2472 | -106.7245
24.2509 | -106.7114
24.2546 | -106.683
24.2582 | -106.6033
24.2619 | -106.4961
24.2656 | -106.3907
24.2692 | -106.2589
24.2729 | -106.1505
24.2766 | -106.0745
24.2803 | -106.0003
24.2839 | -105.9244
24.2876 | -105.8437
24.2913 | -105.757
24.295 | -105.6661
24.2986 | -105.5697
24.3023 | -105.4688
24.306 | -105.3634
24.3097 | -105.2544
24.3133 | -105.1412
24.317 | -105.0241
24.3207 | -104.9031
24.3243 | -104.7796
24.328 | -104.6564
24.3317 | -104.531
24.3354 | -104.4035
24.339 | -104.2751
24.3427 | -104.1465
24.3464 | -104.0184
24.3501 | -103.8906
24.3537 | -103.7641
24.3574 | -103.6399
24.3611 | -103.5202
24.3648 | -103.4023
24.3684 | -103.288
24.3721 | -103.1777
24.3758 | -103.0898
24.3794 | -102.9922
24.3831 | -102.9001
24.3868 | -102.8137
24.3905 | -102.7324
24.3941 | -102.656
24.3978 | -102.5856
24.4015 | -102.5219
24.4052 | -102.4645
24.4088 | -102.4127
24.4125 | -102.3679
24.4162 | -102.331
24.4199 | -102.301
24.4235 | -102.2771
24.4272 | -102.2605
24.4309 | -102.2508
24.4345 | -102.2476
24.4382 | -102.2516
24.4419 | -102.2624
24.4456 | -102.2804
24.4492 | -102.3055
24.4529 | -102.3371
24.4566 | -102.3763
24.4603 | -102.4223
24.4639 | -102.4739
24.4676 | -102.5318
24.4713 | -102.5957
24.475 | -102.6666
24.4786 | -102.7425
24.4823 | -102.8237
24.486 | -102.911
24.4896 | -103.0034
24.4933 | -103.0994
24.497 | -103.1998
24.5007 | -103.3042
24.5043 | -103.4129
24.508 | -103.525
24.5117 | -103.6397
24.5154 | -103.7572
24.519 | -103.8757
24.5227 | -103.9954
24.5264 | -104.1162
24.5301 | -104.2372
24.5337 | -104.3583
24.5374 | -104.4796
24.5411 | -104.6007
24.5447 | -104.7198
24.5484 | -104.8379
24.5521 | -104.9524
24.5558 | -105.0629
24.5594 | -105.1709
24.5631 | -105.2725
24.5668 | -105.3618
24.5705 | -105.4387
24.5741 | -105.5062
24.5778 | -105.5627
24.5815 | -105.6134
24.5852 | -105.6673
24.5888 | -105.7182
24.5925 | -105.767
24.5962 | -105.8203
24.5998 | -105.8695
24.6035 | -105.9395
24.6072 | -106.0109
24.6109 | -106.055
24.6145 | -106.1022
24.6182 | -106.107
24.6219 | -106.0679
24.6256 | -106.0195
24.6292 | -105.9671
24.6329 | -105.913
24.6366 | -105.8579
24.6403 | -105.7995
24.6439 | -105.7376
24.6476 | -105.673
24.6513 | -105.607
24.6549 | -105.538
24.6586 | -105.4658
24.6623 | -105.3935
24.666 | -105.3217
24.6696 | -105.2487
24.6733 | -105.1728
24.677 | -105.098
24.6807 | -105.0246
24.6843 | -104.9538
24.688 | -104.8832
24.6917 | -104.8133
24.6953 | -104.7445
24.699 | -104.6779
24.7027 | -104.6143
24.7064 | -104.5526
24.71 | -104.4928
24.7137 | -104.4357
24.7174 | -104.3817
24.7211 | -104.3319
24.7247 | -104.2851
24.7284 | -104.2419
24.7321 | -104.2014
24.7358 | -104.1642
24.7394 | -104.132
24.7431 | -104.1036
24.7468 | -104.0787
24.7504 | -104.0582
24.7541 | -104.042
24.7578 | -104.031
24.7615 | -104.0238
24.7651 | -104.0188
24.7688 | -104.0172
24.7725 | -104.0215
24.7762 | -104.0306
24.7798 | -104.0436
24.7835 | -104.0597
24.7872 | -104.0806
24.7909 | -104.1057
24.7945 | -104.1345
24.7982 | -104.1665
24.8019 | -104.2025
24.8055 | -104.241
24.8092 | -104.2828
24.8129 | -104.3277
24.8166 | -104.3749
24.8202 | -104.4249
24.8239 | -104.4757
24.8276 | -104.5292
24.8313 | -104.5858
24.8349 | -104.644
24.8386 | -104.7029
24.8423 | -104.7634
24.846 | -104.8254
24.8496 | -104.8901
24.8533 | -104.9545
24.857 | -105.0187
24.8606 | -105.0813
24.8643 | -105.1431
24.868 | -105.2059
24.8717 | -105.2686
24.8753 | -105.331
24.879 | -105.3932
24.8827 | -105.454
24.8864 | -105.5135
24.89 | -105.5714
24.8937 | -105.627
24.8974 | -105.6797
24.9011 | -105.7321
24.9047 | -105.7619
24.9084 | -105.8121
24.9121 | -105.8594
24.9157 | -105.908
24.9194 | -105.9599
24.9231 | -106.0142
24.9268 | -106.0677
24.9304 | -106.1212
24.9341 | -106.1678
24.9378 | -106.199
24.9415 | -106.2223
24.9451 | -106.2408
24.9488 | -106.2553
24.9525 | -106.2745
24.9562 | -106.2947
24.9598 | -106.3146
24.9635 | -106.3345
24.9672 | -106.3494
24.9708 | -106.3624
24.9745 | -106.3697
24.9782 | -106.374
24.9819 | -106.3742
24.9855 | -106.372
24.9892 | -106.3695
24.9929 | -106.3675
24.9966 | -106.3659
25.0002 | -106.3641
25.0039 | -106.3625
25.0076 | -106.3596
25.0112 | -106.357
25.0149 | -106.3533
25.0186 | -106.3478
25.0223 | -106.3447
25.0259 | -106.3433
25.0296 | -106.3387
25.0333 | -106.3319
25.037 | -106.3246
25.0406 | -106.3191
25.0443 | -106.3117
25.048 | -106.305
25.0517 | -106.2973
25.0553 | -106.2889
25.059 | -106.2806
25.0627 | -106.273
25.0663 | -106.2656
25.07 | -106.2573
25.0737 | -106.2493
25.0774 | -106.2427
25.081 | -106.2366
25.0847 | -106.2306
25.0884 | -106.2253
25.0921 | -106.2184
25.0957 | -106.2132
25.0994 | -106.2077
25.1031 | -106.2011
25.1068 | -106.1954
25.1104 | -106.1898
25.1141 | -106.1844
25.1178 | -106.18
25.1214 | -106.1774
25.1251 | -106.1717
25.1288 | -106.1678
25.1325 | -106.1645
25.1361 | -106.1601
25.1398 | -106.1541
25.1435 | -106.1483
25.1472 | -106.1427
25.1508 | -106.1372
25.1545 | -106.1327
25.1582 | -106.1288
25.1619 | -106.1252
25.1655 | -106.1226
25.1692 | -106.119
25.1729 | -106.1157
25.1765 | -106.1139
25.1802 | -106.1107
25.1839 | -106.1063
25.1876 | -106.1002
25.1912 | -106.0943
25.1949 | -106.0902
25.1986 | -106.086
25.2023 | -106.0797
25.2059 | -106.0735
25.2096 | -106.0683
25.2133 | -106.064
25.217 | -106.0618
25.2206 | -106.0585
25.2243 | -106.0544
25.228 | -106.0501
25.2316 | -106.0464
25.2353 | -106.0424
25.239 | -106.0379
25.2427 | -106.0323
25.2463 | -106.0297
25.25 | -106.0273
25.2537 | -106.0239
25.2574 | -106.0209
25.261 | -106.0167
25.2647 | -106.0137
25.2684 | -106.0101
25.2721 | -106.0068
25.2757 | -106.0056
25.2794 | -106.0064
25.2831 | -106.0082
25.2867 | -106.0093
25.2904 | -106.0098
25.2941 | -106.0102
25.2978 | -106.0097
25.3014 | -106.009
25.3051 | -106.0121
25.3088 | -106.0138
25.3125 | -106.016
25.3161 | -106.0169
25.3198 | -106.0177
25.3235 | -106.0184
25.3271 | -106.0209
25.3308 | -106.0253
25.3345 | -106.0292
25.3382 | -106.0331
25.3418 | -106.0379
25.3455 | -106.0433
25.3492 | -106.0499
25.3529 | -106.055
25.3565 | -106.0592
25.3602 | -106.0618
25.3639 | -106.0636
25.3676 | -106.0646
25.3712 | -106.0673
25.3749 | -106.0696
25.3786 | -106.0698
25.3822 | -106.0689
25.3859 | -106.0675
25.3896 | -106.0672
25.3933 | -106.0669
25.3969 | -106.065
25.4006 | -106.063
25.4043 | -106.0606
25.408 | -106.0593
25.4116 | -106.056
25.4153 | -106.0532
25.419 | -106.0505
25.4227 | -106.046
25.4263 | -106.0416
25.43 | -106.0358
25.4337 | -106.0274
25.4373 | -106.02
25.441 | -106.0113
25.4447 | -106.0022
25.4484 | -105.9937
25.452 | -105.9864
25.4557 | -105.9786
25.4594 | -105.9712
25.4631 | -105.9638
25.4667 | -105.9553
25.4704 | -105.946
25.4741 | -105.9382
25.4778 | -105.9295
25.4814 | -105.9192
25.4851 | -105.9106
25.4888 | -105.9012
25.4924 | -105.8927
25.4961 | -105.8856
25.4998 | -105.8776
25.5035 | -105.87
25.5071 | -105.8628
25.5108 | -105.8546
25.5145 | -105.8454
25.5182 | -105.8373
25.5218 | -105.8306
25.5255 | -105.8225
25.5292 | -105.8159
25.5329 | -105.8098
25.5365 | -105.8021
25.5402 | -105.7962
25.5439 | -105.7913
25.5475 | -105.7878
25.5512 | -105.7843
25.5549 | -105.7803
25.5586 | -105.7754
25.5622 | -105.7717
25.5659 | -105.7677
25.5696 | -105.7639
25.5733 | -105.7615
25.5769 | -105.7586
25.5806 | -105.7566
25.5843 | -105.7551
25.588 | -105.7541
25.5916 | -105.7533
25.5953 | -105.7509
25.599 | -105.7496
25.6026 | -105.7508
25.6063 | -105.7519
25.61 | -105.7532
25.6137 | -105.754
25.6173 | -105.7543
25.621 | -105.7558
25.6247 | -105.7573
25.6284 | -105.7605
25.632 | -105.7628
25.6357 | -105.7646
25.6394 | -105.7678
25.6431 | -105.771
25.6467 | -105.7735
25.6504 | -105.7776
25.6541 | -105.7829
25.6577 | -105.7871
25.6614 | -105.7933
25.6651 | -105.7981
25.6688 | -105.8024
25.6724 | -105.8076
25.6761 | -105.814
25.6798 | -105.8212
25.6835 | -105.8291
25.6871 | -105.8345
25.6908 | -105.8387
25.6945 | -105.8447
25.6981 | -105.8517
25.7018 | -105.8588
25.7055 | -105.8658
25.7092 | -105.8729
25.7128 | -105.8819
25.7165 | -105.8912
25.7202 | -105.9008
25.7239 | -105.9125
25.7275 | -105.9254
25.7312 | -105.9354
25.7349 | -105.9437
25.7386 | -105.9536
25.7422 | -105.9637
25.7459 | -105.9742
25.7496 | -105.9843
25.7532 | -105.9944
25.7569 | -106.0058
25.7606 | -106.0176
25.7643 | -106.0297
25.7679 | -106.041
25.7716 | -106.0523
25.7753 | -106.0626
25.779 | -106.0745
25.7826 | -106.0868
25.7863 | -106.0996
25.79 | -106.1127
25.7937 | -106.123
25.7973 | -106.134
25.801 | -106.1464
25.8047 | -106.1603
25.8083 | -106.1755
25.812 | -106.1888
25.8157 | -106.2013
25.8194 | -106.2123
25.823 | -106.2223
25.8267 | -106.2336
25.8304 | -106.2452
25.8341 | -106.2563
25.8377 | -106.2652
25.8414 | -106.2764
25.8451 | -106.2875
25.8488 | -106.2988
25.8524 | -106.3081
25.8561 | -106.3165
25.8598 | -106.3251
25.8634 | -106.3331
25.8671 | -106.3397
25.8708 | -106.3464
25.8745 | -106.355
25.8781 | -106.3639
25.8818 | -106.3702
25.8855 | -106.3736
25.8892 | -106.3803
25.8928 | -106.3868
25.8965 | -106.3918
25.9002 | -106.3969
25.9039 | -106.4
25.9075 | -106.4025
25.9112 | -106.4054
25.9149 | -106.4073
25.9185 | -106.4102
25.9222 | -106.4145
25.9259 | -106.418
25.9296 | -106.4214
25.9332 | -106.4227
25.9369 | -106.4237
25.9406 | -106.4257
25.9443 | -106.428
25.9479 | -106.43
25.9516 | -106.43
25.9553 | -106.4288
25.959 | -106.4305
25.9626 | -106.4314
25.9663 | -106.4312
25.97 | -106.4309
25.9736 | -106.4301
25.9773 | -106.4283
25.981 | -106.4286
25.9847 | -106.4286
25.9883 | -106.4282
25.992 | -106.428
25.9957 | -106.4271
25.9994 | -106.4266
26.003 | -106.4244
26.0067 | -106.4187
26.0104 | -106.4145
26.0141 | -106.4126
26.0177 | -106.4101
26.0214 | -106.4076
26.0251 | -106.4064
26.0287 | -106.4031
26.0324 | -106.4008
26.0361 | -106.3996
26.0398 | -106.397
26.0434 | -106.3958
26.0471 | -106.3936
26.0508 | -106.3901
26.0545 | -106.3882
26.0581 | -106.3861
26.0618 | -106.3842
26.0655 | -106.3823
26.0691 | -106.3785
26.0728 | -106.3748
26.0765 | -106.3716
26.0802 | -106.3682
26.0838 | -106.3669
26.0875 | -106.3655
26.0912 | -106.3635
26.0949 | -106.3607
26.0985 | -106.3585
26.1022 | -106.3585
26.1059 | -106.357
26.1096 | -106.356
26.1132 | -106.3551
26.1169 | -106.3553
26.1206 | -106.3552
26.1242 | -106.3525
26.1279 | -106.3509
26.1316 | -106.3495
26.1353 | -106.3499
26.1389 | -106.3497
26.1426 | -106.3499
26.1463 | -106.3496
26.15 | -106.3501
26.1536 | -106.3514
26.1573 | -106.3539
26.161 | -106.358
26.1647 | -106.3619
26.1683 | -106.3655
26.172 | -106.368
26.1757 | -106.3694
26.1793 | -106.3708
26.183 | -106.3735
26.1867 | -106.3777
26.1904 | -106.3839
26.194 | -106.3905
26.1977 | -106.3952
26.2014 | -106.4008
26.2051 | -106.4074
26.2087 | -106.4146
26.2124 | -106.4226
26.2161 | -106.4301
26.2198 | -106.4372
26.2234 | -106.4451
26.2271 | -106.4546
26.2308 | -106.4653
26.2344 | -106.4752
26.2381 | -106.4834
26.2418 | -106.4934
26.2455 | -106.504
26.2491 | -106.5139
26.2528 | -106.5246
26.2565 | -106.5336
26.2602 | -106.5434
26.2638 | -106.553
26.2675 | -106.5638
26.2712 | -106.5754
26.2749 | -106.5872
26.2785 | -106.5982
26.2822 | -106.6098
26.2859 | -106.6183
26.2895 | -106.6255
26.2932 | -106.6321
26.2969 | -106.6389
26.3006 | -106.6472
26.3042 | -106.6537
26.3079 | -106.6594
26.3116 | -106.6637
26.3153 | -106.6683
26.3189 | -106.6741
26.3226 | -106.6799
26.3263 | -106.6842
26.33 | -106.6883
26.3336 | -106.6913
26.3373 | -106.6936
26.341 | -106.6946
26.3446 | -106.6953
26.3483 | -106.6963
26.352 | -106.696
26.3557 | -106.695
26.3593 | -106.6917
26.363 | -106.6888
26.3667 | -106.6859
26.3704 | -106.6821
26.374 | -106.6789
26.3777 | -106.6783
26.3814 | -106.6744
26.385 | -106.6693
26.3887 | -106.6637
26.3924 | -106.6578
26.3961 | -106.6521
26.3997 | -106.6472
26.4034 | -106.6404
26.4071 | -106.6318
26.4108 | -106.6233
26.4144 | -106.6146
26.4181 | -106.6073
26.4218 | -106.6004
26.4255 | -106.5938
26.4291 | -106.587
26.4328 | -106.5793
26.4365 | -106.5722
26.4401 | -106.5653
26.4438 | -106.5583
26.4475 | -106.5505
26.4512 | -106.5426
26.4548 | -106.5341
26.4585 | -106.5279
26.4622 | -106.5209
26.4659 | -106.5138
26.4695 | -106.5079
26.4732 | -106.5006
26.4769 | -106.4946
26.4806 | -106.4888
26.4842 | -106.4824
26.4879 | -106.4756
26.4916 | -106.4675
26.4952 | -106.4633
26.4989 | -106.4599
26.5026 | -106.4573
26.5063 | -106.4532
26.5099 | -106.4467
26.5136 | -106.4434
26.5173 | -106.4414
26.521 | -106.4379
26.5246 | -106.4345
26.5283 | -106.4308
26.532 | -106.4277
26.5357 | -106.4259
26.5393 | -106.4231
26.543 | -106.4194
26.5467 | -106.4173
26.5503 | -106.4149
26.554 | -106.4139
26.5577 | -106.4125
26.5614 | -106.4104
26.565 | -106.4074
26.5687 | -106.403
26.5724 | -106.4003
26.5761 | -106.3987
26.5797 | -106.3974
26.5834 | -106.3939
26.5871 | -106.3904
26.5908 | -106.3869
26.5944 | -106.3845
26.5981 | -106.3797
26.6018 | -106.3742
26.6054 | -106.3715
26.6091 | -106.3701
26.6128 | -106.3687
26.6165 | -106.3651
26.6201 | -106.3616
26.6238 | -106.36
26.6275 | -106.3559
26.6312 | -106.3522
26.6348 | -106.3474
26.6385 | -106.3444
26.6422 | -106.341
26.6459 | -106.3368
26.6495 | -106.3327
26.6532 | -106.3274
26.6569 | -106.3222
26.6605 | -106.319
26.6642 | -106.3158
26.6679 | -106.3129
26.6716 | -106.3114
26.6752 | -106.3093
26.6789 | -106.3071
26.6826 | -106.3039
26.6863 | -106.3013
26.6899 | -106.2991
26.6936 | -106.2975
26.6973 | -106.2973
26.7009 | -106.2979
26.7046 | -106.2968
26.7083 | -106.296
26.712 | -106.2958
26.7156 | -106.2964
26.7193 | -106.2977
26.723 | -106.3005
26.7267 | -106.3033
26.7303 | -106.3059
26.734 | -106.3092
26.7377 | -106.3134
26.7414 | -106.3182
26.745 | -106.3227
26.7487 | -106.3295
26.7524 | -106.3342
26.756 | -106.3393
26.7597 | -106.3437
26.7634 | -106.3482
26.7671 | -106.3547
26.7707 | -106.3617
26.7744 | -106.3686
26.7781 | -106.3762
26.7818 | -106.3846
26.7854 | -106.3922
26.7891 | -106.3997
26.7928 | -106.4081
26.7965 | -106.416
26.8001 | -106.4229
26.8038 | -106.4306
26.8075 | -106.438
26.8111 | -106.4448
26.8148 | -106.4528
26.8185 | -106.4595
26.8222 | -106.4668
26.8258 | -106.4742
26.8295 | -106.4817
26.8332 | -106.4882
26.8369 | -106.4965
26.8405 | -106.5032
26.8442 | -106.5082
26.8479 | -106.5137
26.8516 | -106.5172
26.8552 | -106.521
26.8589 | -106.5267
26.8626 | -106.5325
26.8662 | -106.5382
26.8699 | -106.5427
26.8736 | -106.5441
26.8773 | -106.5458
26.8809 | -106.5457
26.8846 | -106.5477
26.8883 | -106.5503
26.892 | -106.5492
26.8956 | -106.5466
26.8993 | -106.5462
26.903 | -106.5461
26.9067 | -106.5455
26.9103 | -106.5435
26.914 | -106.5399
26.9177 | -106.5367
26.9213 | -106.5339
26.925 | -106.5303
26.9287 | -106.5259
26.9324 | -106.5205
26.936 | -106.5149
26.9397 | -106.5085
26.9434 | -106.5004
26.9471 | -106.4949
26.9507 | -106.4882
26.9544 | -106.4801
26.9581 | -106.4717
26.9618 | -106.4644
26.9654 | -106.4561
26.9691 | -106.4467
26.9728 | -106.4378
26.9764 | -106.4312
26.9801 | -106.4237
26.9838 | -106.4132
26.9875 | -106.4013
26.9911 | -106.391
26.9948 | -106.3815
26.9985 | -106.3714
27.0022 | -106.3601
27.0058 | -106.3524
27.0095 | -106.3449
27.0132 | -106.3367
27.0169 | -106.3276
27.0205 | -106.3181
27.0242 | -106.3126
27.0279 | -106.3039
27.0315 | -106.2961
27.0352 | -106.2903
27.0389 | -106.2833
27.0426 | -106.2774
27.0462 | -106.2718
27.0499 | -106.2655
27.0536 | -106.2594
27.0573 | -106.2529
27.0609 | -106.2482
27.0646 | -106.2434
27.0683 | -106.2391
27.072 | -106.2354
27.0756 | -106.2307
27.0793 | -106.2265
27.083 | -106.2232
27.0866 | -106.2213
27.0903 | -106.2203
27.094 | -106.2191
27.0977 | -106.2182
27.1013 | -106.216
27.105 | -106.2136
27.1087 | -106.2132
27.1124 | -106.2124
27.116 | -106.2109
27.1197 | -106.2108
27.1234 | -106.21
27.127 | -106.2104
27.1307 | -106.2093
27.1344 | -106.2096
27.1381 | -106.2113
27.1417 | -106.2133
27.1454 | -106.2155
27.1491 | -106.2151
27.1528 | -106.2154
27.1564 | -106.2143
27.1601 | -106.2136
27.1638 | -106.2137
27.1675 | -106.2141
27.1711 | -106.2143
27.1748 | -106.2139
27.1785 | -106.2149
27.1821 | -106.2159
27.1858 | -106.2163
27.1895 | -106.2167
27.1932 | -106.2167
27.1968 | -106.2166
27.2005 | -106.2164
27.2042 | -106.2151
27.2079 | -106.2114
27.2115 | -106.2072
27.2152 | -106.2052
27.2189 | -106.201
27.2226 | -106.1978
27.2262 | -106.1956
27.2299 | -106.1934
27.2336 | -106.1908
27.2372 | -106.1875
27.2409 | -106.1828
27.2446 | -106.1777
27.2483 | -106.1734
27.2519 | -106.1701
27.2556 | -106.1651
27.2593 | -106.1603
27.263 | -106.1569
27.2666 | -106.1515
27.2703 | -106.1468
27.274 | -106.1427
27.2777 | -106.1386
27.2813 | -106.133
27.285 | -106.1261
27.2887 | -106.1192
27.2923 | -106.1137
27.296 | -106.1091
27.2997 | -106.1046
27.3034 | -106.1005
27.307 | -106.0943
27.3107 | -106.0895
27.3144 | -106.0848
27.3181 | -106.0787
27.3217 | -106.0743
27.3254 | -106.0717
27.3291 | -106.0662
27.3328 | -106.0604
27.3364 | -106.0557
27.3401 | -106.0511
27.3438 | -106.0465
27.3474 | -106.043
27.3511 | -106.0411
27.3548 | -106.0376
27.3585 | -106.0339
27.3621 | -106.0305
27.3658 | -106.0281
27.3695 | -106.0257
27.3732 | -106.0231
27.3768 | -106.0194
27.3805 | -106.0166
27.3842 | -106.0146
27.3879 | -106.0116
27.3915 | -106.0083
27.3952 | -106.0074
27.3989 | -106.0061
27.4025 | -106.0049
27.4062 | -106.0034
27.4099 | -106.0027
27.4136 | -106.003
27.4172 | -106.0009
27.4209 | -105.9993
27.4246 | -105.9999
27.4283 | -105.9985
27.4319 | -105.9973
27.4356 | -105.9981
27.4393 | -105.9985
27.4429 | -105.9988
27.4466 | -105.9991
27.4503 | -105.9982
27.454 | -105.9976
27.4576 | -105.9982
27.4613 | -105.9988
27.465 | -105.9978
27.4687 | -105.9972
27.4723 | -105.9976
27.476 | -105.9974
27.4797 | -105.9985
27.4834 | -105.9985
27.487 | -105.9996
27.4907 | -106.0002
27.4944 | -105.9989
27.498 | -105.9986
27.5017 | -105.9999
27.5054 | -105.9999
27.5091 | -106.0001
27.5127 | -105.9998
27.5164 | -105.9997
27.5201 | -106.0005
27.5238 | -106.0013
27.5274 | -106.0009
27.5311 | -106.0002
27.5348 | -106.0008
27.5385 | -106.0008
27.5421 | -106.0005
27.5458 | -105.9993
27.5495 | -105.9976
27.5531 | -105.9963
27.5568 | -105.996
27.5605 | -105.9953
27.5642 | -105.9936
27.5678 | -105.9899
27.5715 | -105.9878
27.5752 | -105.9863
27.5789 | -105.9853
27.5825 | -105.9849
27.5862 | -105.9839
27.5899 | -105.9817
27.5936 | -105.9808
27.5972 | -105.9813
27.6009 | -105.9797
27.6046 | -105.9762
27.6082 | -105.9739
27.6119 | -105.9709
27.6156 | -105.9678
27.6193 | -105.9651
27.6229 | -105.9611
27.6266 | -105.956
27.6303 | -105.9522
27.634 | -105.9493
27.6376 | -105.9468
27.6413 | -105.9428
27.645 | -105.9395
27.6487 | -105.9348
27.6523 | -105.9288
27.656 | -105.9236
27.6597 | -105.9183
27.6633 | -105.9147
27.667 | -105.91
27.6707 | -105.9049
27.6744 | -105.9
27.678 | -105.897
27.6817 | -105.8915
27.6854 | -105.8842
27.6891 | -105.878
27.6927 | -105.872
27.6964 | -105.8665
27.7001 | -105.862
27.7038 | -105.8572
27.7074 | -105.8531
27.7111 | -105.8478
27.7148 | -105.8415
27.7184 | -105.8354
27.7221 | -105.8301
27.7258 | -105.824
27.7295 | -105.8187
27.7331 | -105.8139
27.7368 | -105.808
27.7405 | -105.8005
27.7442 | -105.794
27.7478 | -105.7893
27.7515 | -105.7848
27.7552 | -105.781
27.7588 | -105.7775
27.7625 | -105.7746
27.7662 | -105.7721
27.7699 | -105.7691
27.7735 | -105.7661
27.7772 | -105.763
27.7809 | -105.761
27.7846 | -105.7572
27.7882 | -105.7546
27.7919 | -105.7538
27.7956 | -105.7526
27.7993 | -105.75
27.8029 | -105.7481
27.8066 | -105.7491
27.8103 | -105.7499
27.8139 | -105.7495
27.8176 | -105.7504
27.8213 | -105.7525
27.825 | -105.7538
27.8286 | -105.7548
27.8323 | -105.7571
27.836 | -105.7607
27.8397 | -105.7654
27.8433 | -105.7698
27.847 | -105.7735
27.8507 | -105.7771
27.8544 | -105.7794
27.858 | -105.78
27.8617 | -105.7813
27.8654 | -105.7848
27.869 | -105.7871
27.8727 | -105.7912
27.8764 | -105.7964
27.8801 | -105.8001
27.8837 | -105.8041
27.8874 | -105.8101
27.8911 | -105.8153
27.8948 | -105.818
27.8984 | -105.8202
27.9021 | -105.8226
27.9058 | -105.8248
27.9095 | -105.8265
27.9131 | -105.8288
27.9168 | -105.8311
27.9205 | -105.8321
27.9241 | -105.8234
27.9278 | -105.8029
27.9315 | -105.7846
27.9352 | -105.7728
27.9388 | -105.766
27.9425 | -105.7577
27.9462 | -105.7483
27.9499 | -105.7382
27.9535 | -105.7268
27.9572 | -105.716
27.9609 | -105.7051
27.9646 | -105.6899
27.9682 | -105.6732
27.9719 | -105.6561
27.9756 | -105.6375
27.9792 | -105.6181
27.9829 | -105.6009
27.9866 | -105.5814
27.9903 | -105.5601
27.9939 | -105.5376
27.9976 | -105.5138
28.0013 | -105.4891