## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 284 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 287
Hood | 02-92B/P | 221
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1012 | hPa |  | PRC PRO 01
Humidity | 54.8 | %/Hr
Internal temperature of calibration bench | 24.14 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1019 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.6 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.1 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3107 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 05/09/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 10/09/2018
Technical Manager | Anthony HUE | Approval |  | 27/09/2018
CEO | Robert GLEMOT | Approval |  | 27/09/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3169 | 57.4323
0.012209 | -35.8891 | 38.2361
0.018313 | -34.8274 | 28.116
0.024417 | -34.3762 | 21.8192
0.030521 | -34.2646 | 18.0915
0.036626 | -34.1226 | 15.0398
0.04273 | -34.0278 | 12.6589
0.048834 | -33.9839 | 11.1302
0.054938 | -33.9224 | 9.9944
0.061043 | -33.8629 | 8.9288
0.067147 | -33.8522 | 8.1051
0.073251 | -33.8535 | 7.3351
0.079355 | -33.872 | 6.6641
0.08546 | -33.8862 | 6.0716
0.091564 | -33.8822 | 5.6373
0.097668 | -33.8655 | 5.3922
0.10377 | -33.8444 | 5.1432
0.10988 | -33.8278 | 4.8656
0.11598 | -33.8188 | 4.5731
0.12209 | -33.8209 | 4.2909
0.12819 | -33.8312 | 4.0094
0.13429 | -33.8425 | 3.7227
0.1404 | -33.8454 | 3.4466
0.1465 | -33.8471 | 3.1915
0.15261 | -33.8443 | 2.9736
0.15871 | -33.838 | 2.8122
0.16482 | -33.8306 | 2.7063
0.17092 | -33.8294 | 2.6603
0.17702 | -33.8281 | 2.5833
0.18313 | -33.8278 | 2.5015
0.18923 | -33.8278 | 2.391
0.19534 | -33.8267 | 2.2647
0.20144 | -33.8256 | 2.1241
0.20754 | -33.8231 | 2.0055
0.21365 | -33.8225 | 1.8925
0.21975 | -33.8224 | 1.7955
0.22586 | -33.824 | 1.7087
0.23196 | -33.8269 | 1.6361
0.23807 | -33.8313 | 1.5726
0.24417 | -33.8354 | 1.5118
0.25027 | -33.8395 | 1.448
0.25638 | -33.842 | 1.3844
0.26248 | -33.8439 | 1.3166
0.26859 | -33.8451 | 1.2506
0.27469 | -33.8458 | 1.1795
0.2808 | -33.8473 | 1.1077
0.2869 | -33.8486 | 1.0381
0.293 | -33.85 | 0.97239
0.29911 | -33.8513 | 0.91381
0.30521 | -33.8526 | 0.85603
0.31132 | -33.8536 | 0.81828
0.31742 | -33.8551 | 0.77616
0.32353 | -33.8557 | 0.74182
0.32963 | -33.8568 | 0.71336
0.33573 | -33.8581 | 0.69307
0.34184 | -33.8596 | 0.67155
0.34794 | -33.8619 | 0.64688
0.35405 | -33.8642 | 0.6221
0.36015 | -33.8668 | 0.59592
0.36626 | -33.8689 | 0.56727
0.37236 | -33.8712 | 0.52957
0.37846 | -33.8734 | 0.48444
0.38457 | -33.8744 | 0.42724
0.39067 | -33.8756 | 0.38005
0.39678 | -33.8764 | 0.32801
0.40288 | -33.877 | 0.27401
0.40899 | -33.8771 | 0.22235
0.41509 | -33.8767 | 0.1719
0.42119 | -33.877 | 0.12928
0.4273 | -33.8772 | 0.08345
0.4334 | -33.8774 | 0.037669
0.43951 | -33.8779 | -0.0053972
0.44561 | -33.879 | -0.045121
0.45172 | -33.8808 | -0.0769
0.45782 | -33.8818 | -0.11395
0.46392 | -33.8833 | -0.15019
0.47003 | -33.8846 | -0.1825
0.47613 | -33.8856 | -0.21661
0.48224 | -33.8863 | -0.25147
0.48834 | -33.8855 | -0.28656
0.49445 | -33.8855 | -0.32202
0.50055 | -33.8851 | -0.35448
0.50665 | -33.8845 | -0.38918
0.51276 | -33.8836 | -0.42371
0.51886 | -33.8827 | -0.45815
0.52497 | -33.8815 | -0.48306
0.53107 | -33.8809 | -0.51319
0.53717 | -33.8803 | -0.54204
0.54328 | -33.8795 | -0.56577
0.54938 | -33.8794 | -0.58903
0.55549 | -33.8795 | -0.6107
0.56159 | -33.8807 | -0.62285
0.5677 | -33.8811 | -0.63802
0.5738 | -33.8823 | -0.65487
0.5799 | -33.8831 | -0.67033
0.58601 | -33.8843 | -0.68099
0.59211 | -33.886 | -0.68776
0.59822 | -33.8874 | -0.69774
0.60432 | -33.889 | -0.71031
0.61043 | -33.8903 | -0.7187
0.61653 | -33.8916 | -0.7293
0.62263 | -33.8932 | -0.73892
0.62874 | -33.8939 | -0.74971
0.63484 | -33.8948 | -0.75915
0.64095 | -33.8953 | -0.76956
0.64705 | -33.8958 | -0.77946
0.65316 | -33.896 | -0.79211
0.65926 | -33.8962 | -0.80808
0.66536 | -33.8951 | -0.82575
0.67147 | -33.8948 | -0.83915
0.67757 | -33.8945 | -0.85369
0.68368 | -33.8939 | -0.87275
0.68978 | -33.8932 | -0.89005
0.69589 | -33.8924 | -0.90923
0.70199 | -33.8916 | -0.93223
0.70809 | -33.8908 | -0.95159
0.7142 | -33.8903 | -0.97218
0.7203 | -33.8893 | -0.99161
0.72641 | -33.8888 | -1.0143
0.73251 | -33.8883 | -1.0345
0.73862 | -33.8887 | -1.0565
0.74472 | -33.8886 | -1.0768
0.75082 | -33.8885 | -1.0991
0.75693 | -33.889 | -1.1212
0.76303 | -33.8897 | -1.1433
0.76914 | -33.8916 | -1.1639
0.77524 | -33.8927 | -1.1837
0.78135 | -33.8941 | -1.204
0.78745 | -33.8956 | -1.2241
0.79355 | -33.897 | -1.2441
0.79966 | -33.8988 | -1.2637
0.80576 | -33.9012 | -1.2837
0.81187 | -33.9029 | -1.3055
0.81797 | -33.905 | -1.327
0.82408 | -33.9067 | -1.3477
0.83018 | -33.9086 | -1.3685
0.83628 | -33.9106 | -1.3872
0.84239 | -33.9118 | -1.408
0.84849 | -33.9133 | -1.43
0.8546 | -33.9149 | -1.4499
0.8607 | -33.9162 | -1.4708
0.86681 | -33.9173 | -1.49
0.87291 | -33.9184 | -1.509
0.87901 | -33.9189 | -1.5272
0.88512 | -33.9197 | -1.5457
0.89122 | -33.9205 | -1.5642
0.89733 | -33.921 | -1.5801
0.90343 | -33.9215 | -1.5978
0.90953 | -33.9218 | -1.6123
0.91564 | -33.9219 | -1.6283
0.92174 | -33.9222 | -1.6435
0.92785 | -33.9224 | -1.6583
0.93395 | -33.9225 | -1.6723
0.94006 | -33.9228 | -1.6853
0.94616 | -33.9226 | -1.6972
0.95226 | -33.9228 | -1.71
0.95837 | -33.923 | -1.724
0.96447 | -33.9233 | -1.7369
0.97058 | -33.9233 | -1.7503
0.97668 | -33.9231 | -1.7638
0.98279 | -33.923 | -1.7766
0.98889 | -33.9229 | -1.7874
0.99499 | -33.9229 | -1.801
1.0011 | -33.9229 | -1.8118
1.0072 | -33.9226 | -1.8235
1.0133 | -33.9224 | -1.8344
1.0194 | -33.9221 | -1.8476
1.0255 | -33.9219 | -1.8614
1.0316 | -33.9214 | -1.8735
1.0377 | -33.9211 | -1.8873
1.0438 | -33.9206 | -1.9001
1.0499 | -33.9199 | -1.9156
1.056 | -33.9193 | -1.93
1.0621 | -33.9188 | -1.9445
1.0682 | -33.9184 | -1.9571
1.0743 | -33.9178 | -1.9727
1.0805 | -33.9174 | -1.9873
1.0866 | -33.9165 | -2.0026
1.0927 | -33.9158 | -2.0173
1.0988 | -33.9152 | -2.0323
1.1049 | -33.9145 | -2.0467
1.111 | -33.9138 | -2.0623
1.1171 | -33.9132 | -2.077
1.1232 | -33.9128 | -2.0924
1.1293 | -33.9124 | -2.1066
1.1354 | -33.9122 | -2.1219
1.1415 | -33.9119 | -2.1376
1.1476 | -33.9113 | -2.1533
1.1537 | -33.9109 | -2.1681
1.1598 | -33.911 | -2.1832
1.1659 | -33.9109 | -2.1967
1.172 | -33.9107 | -2.2109
1.1781 | -33.9105 | -2.225
1.1842 | -33.9104 | -2.2394
1.1903 | -33.9108 | -2.2506
1.1964 | -33.9107 | -2.2641
1.2025 | -33.9108 | -2.2783
1.2086 | -33.9112 | -2.2909
1.2147 | -33.9113 | -2.3036
1.2209 | -33.9117 | -2.315
1.227 | -33.9124 | -2.3259
1.2331 | -33.9129 | -2.338
1.2392 | -33.9133 | -2.35
1.2453 | -33.9137 | -2.3612
1.2514 | -33.9142 | -2.3712
1.2575 | -33.9148 | -2.3823
1.2636 | -33.9156 | -2.3912
1.2697 | -33.9163 | -2.4006
1.2758 | -33.917 | -2.4117
1.2819 | -33.9177 | -2.4217
1.288 | -33.9184 | -2.4325
1.2941 | -33.9193 | -2.4411
1.3002 | -33.92 | -2.4492
1.3063 | -33.9206 | -2.4583
1.3124 | -33.9213 | -2.4673
1.3185 | -33.922 | -2.4764
1.3246 | -33.9228 | -2.4855
1.3307 | -33.9236 | -2.4941
1.3368 | -33.9245 | -2.5029
1.3429 | -33.9252 | -2.5118
1.349 | -33.9257 | -2.5212
1.3551 | -33.9264 | -2.5305
1.3613 | -33.927 | -2.5388
1.3674 | -33.9277 | -2.5489
1.3735 | -33.9282 | -2.5587
1.3796 | -33.9287 | -2.5678
1.3857 | -33.9291 | -2.5771
1.3918 | -33.9297 | -2.5874
1.3979 | -33.9303 | -2.5967
1.404 | -33.9308 | -2.6082
1.4101 | -33.9313 | -2.6187
1.4162 | -33.9317 | -2.6294
1.4223 | -33.9323 | -2.6409
1.4284 | -33.9328 | -2.6511
1.4345 | -33.9331 | -2.6627
1.4406 | -33.9334 | -2.6764
1.4467 | -33.9339 | -2.6882
1.4528 | -33.9344 | -2.6997
1.4589 | -33.935 | -2.7118
1.465 | -33.9353 | -2.7228
1.4711 | -33.9358 | -2.7344
1.4772 | -33.9364 | -2.748
1.4833 | -33.9369 | -2.7612
1.4894 | -33.9375 | -2.7736
1.4955 | -33.9381 | -2.7858
1.5016 | -33.9387 | -2.7985
1.5078 | -33.9391 | -2.8127
1.5139 | -33.9396 | -2.8259
1.52 | -33.9401 | -2.8388
1.5261 | -33.9405 | -2.8517
1.5322 | -33.9408 | -2.8659
1.5383 | -33.9413 | -2.8788
1.5444 | -33.9419 | -2.8917
1.5505 | -33.9424 | -2.9053
1.5566 | -33.943 | -2.9186
1.5627 | -33.9435 | -2.9326
1.5688 | -33.9442 | -2.9465
1.5749 | -33.9446 | -2.9594
1.581 | -33.9449 | -2.9731
1.5871 | -33.9454 | -2.9871
1.5932 | -33.9458 | -3.0003
1.5993 | -33.9463 | -3.0147
1.6054 | -33.9467 | -3.0283
1.6115 | -33.9472 | -3.0427
1.6176 | -33.9475 | -3.0555
1.6237 | -33.9478 | -3.0695
1.6298 | -33.9481 | -3.0823
1.6359 | -33.9486 | -3.0954
1.642 | -33.949 | -3.1077
1.6482 | -33.9493 | -3.121
1.6543 | -33.9497 | -3.135
1.6604 | -33.9501 | -3.1473
1.6665 | -33.9502 | -3.1601
1.6726 | -33.9505 | -3.173
1.6787 | -33.9509 | -3.1859
1.6848 | -33.9509 | -3.1989
1.6909 | -33.9512 | -3.2119
1.697 | -33.9515 | -3.2245
1.7031 | -33.9516 | -3.2374
1.7092 | -33.9516 | -3.2499
1.7153 | -33.9519 | -3.2622
1.7214 | -33.9517 | -3.2747
1.7275 | -33.9517 | -3.2876
1.7336 | -33.9517 | -3.3009
1.7397 | -33.952 | -3.3135
1.7458 | -33.952 | -3.326
1.7519 | -33.9521 | -3.3384
1.758 | -33.9522 | -3.3507
1.7641 | -33.9524 | -3.3626
1.7702 | -33.9524 | -3.3758
1.7763 | -33.9524 | -3.3876
1.7824 | -33.9525 | -3.4003
1.7885 | -33.9525 | -3.413
1.7947 | -33.9527 | -3.4252
1.8008 | -33.9527 | -3.4368
1.8069 | -33.9528 | -3.4486
1.813 | -33.9529 | -3.461
1.8191 | -33.953 | -3.474
1.8252 | -33.9531 | -3.4869
1.8313 | -33.9533 | -3.4993
1.8374 | -33.9535 | -3.5113
1.8435 | -33.9536 | -3.5234
1.8496 | -33.9537 | -3.5366
1.8557 | -33.954 | -3.5495
1.8618 | -33.9543 | -3.5618
1.8679 | -33.9545 | -3.574
1.874 | -33.9546 | -3.5861
1.8801 | -33.9549 | -3.5986
1.8862 | -33.9552 | -3.6108
1.8923 | -33.9553 | -3.623
1.8984 | -33.9555 | -3.6352
1.9045 | -33.9556 | -3.6468
1.9106 | -33.9557 | -3.6591
1.9167 | -33.9559 | -3.6714
1.9228 | -33.956 | -3.6834
1.9289 | -33.9562 | -3.6953
1.9351 | -33.9564 | -3.7062
1.9412 | -33.9567 | -3.7176
1.9473 | -33.9571 | -3.7298
1.9534 | -33.9573 | -3.7413
1.9595 | -33.9575 | -3.7532
1.9656 | -33.9578 | -3.7655
1.9717 | -33.9581 | -3.7772
1.9778 | -33.9585 | -3.7887
1.9839 | -33.9588 | -3.8003
1.99 | -33.959 | -3.8114
1.9961 | -33.9591 | -3.8236
2.0022 | -33.9594 | -3.8351
2.0083 | -33.9595 | -3.8464
2.0144 | -33.9598 | -3.858
2.0205 | -33.9602 | -3.8698
2.0266 | -33.9605 | -3.8812
2.0327 | -33.9608 | -3.8929
2.0388 | -33.9613 | -3.9045
2.0449 | -33.9615 | -3.916
2.051 | -33.9618 | -3.9275
2.0571 | -33.9621 | -3.9389
2.0632 | -33.9624 | -3.9505
2.0693 | -33.9626 | -3.9628
2.0754 | -33.9629 | -3.9743
2.0816 | -33.9632 | -3.9865
2.0877 | -33.9635 | -3.9982
2.0938 | -33.9639 | -4.0104
2.0999 | -33.9641 | -4.0222
2.106 | -33.9643 | -4.0339
2.1121 | -33.9645 | -4.0461
2.1182 | -33.9648 | -4.057
2.1243 | -33.9651 | -4.0679
2.1304 | -33.9653 | -4.0792
2.1365 | -33.9656 | -4.091
2.1426 | -33.9659 | -4.1032
2.1487 | -33.9662 | -4.1146
2.1548 | -33.9666 | -4.1265
2.1609 | -33.9671 | -4.1381
2.167 | -33.9675 | -4.1495
2.1731 | -33.9678 | -4.1616
2.1792 | -33.968 | -4.1728
2.1853 | -33.9684 | -4.1849
2.1914 | -33.9688 | -4.1964
2.1975 | -33.9692 | -4.208
2.2036 | -33.9694 | -4.2203
2.2097 | -33.9697 | -4.2318
2.2158 | -33.9701 | -4.2426
2.222 | -33.9703 | -4.2532
2.2281 | -33.9706 | -4.2648
2.2342 | -33.9709 | -4.2762
2.2403 | -33.9711 | -4.2871
2.2464 | -33.9714 | -4.298
2.2525 | -33.9717 | -4.3096
2.2586 | -33.9721 | -4.3203
2.2647 | -33.9724 | -4.3309
2.2708 | -33.9727 | -4.3423
2.2769 | -33.9731 | -4.3531
2.283 | -33.9732 | -4.3644
2.2891 | -33.9735 | -4.3755
2.2952 | -33.9737 | -4.3862
2.3013 | -33.9739 | -4.3973
2.3074 | -33.9742 | -4.4079
2.3135 | -33.9744 | -4.4194
2.3196 | -33.9747 | -4.4298
2.3257 | -33.9749 | -4.4406
2.3318 | -33.975 | -4.4517
2.3379 | -33.9753 | -4.4627
2.344 | -33.9754 | -4.4742
2.3501 | -33.9757 | -4.4851
2.3562 | -33.9758 | -4.4956
2.3623 | -33.976 | -4.5062
2.3685 | -33.9761 | -4.5169
2.3746 | -33.9763 | -4.5283
2.3807 | -33.9765 | -4.5389
2.3868 | -33.9767 | -4.5498
2.3929 | -33.9769 | -4.5606
2.399 | -33.977 | -4.5712
2.4051 | -33.9773 | -4.5819
2.4112 | -33.9775 | -4.5927
2.4173 | -33.9777 | -4.6035
2.4234 | -33.9778 | -4.6144
2.4295 | -33.978 | -4.6254
2.4356 | -33.9782 | -4.6364
2.4417 | -33.9784 | -4.6477
2.4478 | -33.9785 | -4.6588
2.4539 | -33.9787 | -4.6692
2.46 | -33.9789 | -4.6808
2.4661 | -33.979 | -4.6921
2.4722 | -33.9792 | -4.7024
2.4783 | -33.9794 | -4.713
2.4844 | -33.9795 | -4.7242
2.4905 | -33.9796 | -4.7354
2.4966 | -33.9798 | -4.7462
2.5027 | -33.9799 | -4.7569
2.5089 | -33.9801 | -4.7679
2.515 | -33.9802 | -4.7789
2.5211 | -33.9803 | -4.7898
2.5272 | -33.9804 | -4.8009
2.5333 | -33.9806 | -4.812
2.5394 | -33.9808 | -4.8229
2.5455 | -33.981 | -4.8344
2.5516 | -33.9812 | -4.845
2.5577 | -33.9814 | -4.8566
2.5638 | -33.9816 | -4.8685
2.5699 | -33.9816 | -4.8801
2.576 | -33.9817 | -4.8918
2.5821 | -33.9819 | -4.9033
2.5882 | -33.9821 | -4.9146
2.5943 | -33.9822 | -4.9256
2.6004 | -33.9825 | -4.9367
2.6065 | -33.9826 | -4.9473
2.6126 | -33.9828 | -4.958
2.6187 | -33.9829 | -4.9686
2.6248 | -33.9831 | -4.9796
2.6309 | -33.9833 | -4.9912
2.637 | -33.9836 | -5.003
2.6431 | -33.9838 | -5.0137
2.6492 | -33.9839 | -5.0253
2.6554 | -33.9842 | -5.0363
2.6615 | -33.9843 | -5.0475
2.6676 | -33.9846 | -5.0591
2.6737 | -33.9849 | -5.0704
2.6798 | -33.9853 | -5.0817
2.6859 | -33.9855 | -5.0932
2.692 | -33.9857 | -5.1043
2.6981 | -33.9859 | -5.1156
2.7042 | -33.9862 | -5.1269
2.7103 | -33.9864 | -5.1377
2.7164 | -33.9868 | -5.1489
2.7225 | -33.9871 | -5.1605
2.7286 | -33.9873 | -5.1717
2.7347 | -33.9876 | -5.1824
2.7408 | -33.9879 | -5.1927
2.7469 | -33.9881 | -5.2041
2.753 | -33.9885 | -5.2149
2.7591 | -33.9888 | -5.2264
2.7652 | -33.9891 | -5.2381
2.7713 | -33.9893 | -5.2491
2.7774 | -33.9896 | -5.2602
2.7835 | -33.9898 | -5.2714
2.7896 | -33.99 | -5.2821
2.7958 | -33.9903 | -5.293
2.8019 | -33.9907 | -5.3036
2.808 | -33.991 | -5.3146
2.8141 | -33.9913 | -5.326
2.8202 | -33.9917 | -5.337
2.8263 | -33.992 | -5.3481
2.8324 | -33.9924 | -5.3597
2.8385 | -33.9928 | -5.3708
2.8446 | -33.9932 | -5.3817
2.8507 | -33.9936 | -5.3927
2.8568 | -33.994 | -5.4034
2.8629 | -33.9943 | -5.4141
2.869 | -33.9946 | -5.4254
2.8751 | -33.995 | -5.4363
2.8812 | -33.9954 | -5.4477
2.8873 | -33.9958 | -5.4588
2.8934 | -33.9961 | -5.47
2.8995 | -33.9965 | -5.4813
2.9056 | -33.9968 | -5.4926
2.9117 | -33.9972 | -5.5038
2.9178 | -33.9976 | -5.5147
2.9239 | -33.998 | -5.5264
2.93 | -33.9984 | -5.5381
2.9361 | -33.9987 | -5.549
2.9423 | -33.9991 | -5.5598
2.9484 | -33.9995 | -5.5711
2.9545 | -33.9999 | -5.5822
2.9606 | -34.0004 | -5.5938
2.9667 | -34.0008 | -5.605
2.9728 | -34.0011 | -5.6165
2.9789 | -34.0016 | -5.6277
2.985 | -34.002 | -5.6391
2.9911 | -34.0024 | -5.6505
2.9972 | -34.0029 | -5.6617
3.0033 | -34.0033 | -5.6734
3.0094 | -34.0037 | -5.6846
3.0155 | -34.0042 | -5.6957
3.0216 | -34.0046 | -5.707
3.0277 | -34.0049 | -5.7189
3.0338 | -34.0053 | -5.7304
3.0399 | -34.0058 | -5.7418
3.046 | -34.0061 | -5.7535
3.0521 | -34.0065 | -5.7647
3.0582 | -34.0068 | -5.7768
3.0643 | -34.0072 | -5.7883
3.0704 | -34.0075 | -5.8003
3.0765 | -34.0078 | -5.8115
3.0827 | -34.0081 | -5.8228
3.0888 | -34.0085 | -5.8342
3.0949 | -34.0089 | -5.8456
3.101 | -34.0093 | -5.8572
3.1071 | -34.0096 | -5.869
3.1132 | -34.01 | -5.8809
3.1193 | -34.0104 | -5.8924
3.1254 | -34.0108 | -5.9039
3.1315 | -34.0112 | -5.9159
3.1376 | -34.0115 | -5.9274
3.1437 | -34.0119 | -5.9391
3.1498 | -34.0122 | -5.9512
3.1559 | -34.0125 | -5.9628
3.162 | -34.0128 | -5.9749
3.1681 | -34.0131 | -5.9869
3.1742 | -34.0135 | -5.9985
3.1803 | -34.0138 | -6.0103
3.1864 | -34.0142 | -6.0216
3.1925 | -34.0145 | -6.0332
3.1986 | -34.0148 | -6.0452
3.2047 | -34.0151 | -6.0574
3.2108 | -34.0153 | -6.0686
3.2169 | -34.0157 | -6.0803
3.223 | -34.0159 | -6.0921
3.2292 | -34.0162 | -6.104
3.2353 | -34.0164 | -6.1149
3.2414 | -34.0167 | -6.1264
3.2475 | -34.0169 | -6.1379
3.2536 | -34.0172 | -6.1494
3.2597 | -34.0175 | -6.1604
3.2658 | -34.0177 | -6.1715
3.2719 | -34.018 | -6.1827
3.278 | -34.0183 | -6.1943
3.2841 | -34.0185 | -6.2055
3.2902 | -34.0188 | -6.217
3.2963 | -34.0191 | -6.2286
3.3024 | -34.0192 | -6.2397
3.3085 | -34.0195 | -6.2514
3.3146 | -34.0197 | -6.2626
3.3207 | -34.02 | -6.2742
3.3268 | -34.0202 | -6.2858
3.3329 | -34.0204 | -6.2966
3.339 | -34.0206 | -6.3074
3.3451 | -34.0208 | -6.3187
3.3512 | -34.021 | -6.3296
3.3573 | -34.0213 | -6.3409
3.3634 | -34.0215 | -6.3521
3.3696 | -34.0217 | -6.3636
3.3757 | -34.0219 | -6.3744
3.3818 | -34.0221 | -6.3854
3.3879 | -34.0223 | -6.3966
3.394 | -34.0225 | -6.4081
3.4001 | -34.0227 | -6.419
3.4062 | -34.0228 | -6.4301
3.4123 | -34.023 | -6.4408
3.4184 | -34.0233 | -6.4522
3.4245 | -34.0235 | -6.4631
3.4306 | -34.0236 | -6.4739
3.4367 | -34.0238 | -6.4846
3.4428 | -34.024 | -6.4951
3.4489 | -34.0242 | -6.5062
3.455 | -34.0244 | -6.5171
3.4611 | -34.0246 | -6.5285
3.4672 | -34.0248 | -6.5397
3.4733 | -34.025 | -6.5509
3.4794 | -34.0252 | -6.5619
3.4855 | -34.0254 | -6.5724
3.4916 | -34.0256 | -6.5829
3.4977 | -34.0257 | -6.5944
3.5038 | -34.0259 | -6.6053
3.5099 | -34.0261 | -6.6158
3.5161 | -34.0263 | -6.6263
3.5222 | -34.0264 | -6.6369
3.5283 | -34.0265 | -6.6478
3.5344 | -34.0267 | -6.6588
3.5405 | -34.0269 | -6.6695
3.5466 | -34.0271 | -6.6805
3.5527 | -34.0272 | -6.6916
3.5588 | -34.0274 | -6.7023
3.5649 | -34.0275 | -6.7132
3.571 | -34.0276 | -6.724
3.5771 | -34.0278 | -6.7352
3.5832 | -34.028 | -6.7453
3.5893 | -34.0282 | -6.756
3.5954 | -34.0284 | -6.767
3.6015 | -34.0286 | -6.7781
3.6076 | -34.0289 | -6.7888
3.6137 | -34.029 | -6.7991
3.6198 | -34.0293 | -6.8095
3.6259 | -34.0295 | -6.8206
3.632 | -34.0297 | -6.8314
3.6381 | -34.0299 | -6.8424
3.6442 | -34.0301 | -6.8531
3.6503 | -34.0303 | -6.8635
3.6565 | -34.0306 | -6.8745
3.6626 | -34.0307 | -6.8852
3.6687 | -34.031 | -6.8961
3.6748 | -34.0313 | -6.9066
3.6809 | -34.0315 | -6.9175
3.687 | -34.0317 | -6.9286
3.6931 | -34.032 | -6.939
3.6992 | -34.0323 | -6.9499
3.7053 | -34.0325 | -6.9606
3.7114 | -34.0326 | -6.9715
3.7175 | -34.0329 | -6.9824
3.7236 | -34.0332 | -6.9927
3.7297 | -34.0334 | -7.0036
3.7358 | -34.0337 | -7.0147
3.7419 | -34.0339 | -7.0257
3.748 | -34.0341 | -7.0365
3.7541 | -34.0344 | -7.0473
3.7602 | -34.0347 | -7.058
3.7663 | -34.0349 | -7.0687
3.7724 | -34.0352 | -7.0796
3.7785 | -34.0355 | -7.0902
3.7846 | -34.0357 | -7.1009
3.7907 | -34.036 | -7.112
3.7969 | -34.0363 | -7.1229
3.803 | -34.0365 | -7.1333
3.8091 | -34.0368 | -7.1441
3.8152 | -34.0371 | -7.1549
3.8213 | -34.0373 | -7.1659
3.8274 | -34.0376 | -7.1764
3.8335 | -34.038 | -7.1874
3.8396 | -34.0383 | -7.1983
3.8457 | -34.0386 | -7.2092
3.8518 | -34.039 | -7.2199
3.8579 | -34.0393 | -7.2307
3.864 | -34.0396 | -7.2417
3.8701 | -34.0399 | -7.2529
3.8762 | -34.0402 | -7.2636
3.8823 | -34.0405 | -7.2746
3.8884 | -34.0408 | -7.2855
3.8945 | -34.0412 | -7.2966
3.9006 | -34.0415 | -7.3078
3.9067 | -34.0418 | -7.3191
3.9128 | -34.0421 | -7.33
3.9189 | -34.0424 | -7.3408
3.925 | -34.0426 | -7.3518
3.9311 | -34.0429 | -7.3622
3.9372 | -34.0432 | -7.3733
3.9434 | -34.0436 | -7.3842
3.9495 | -34.0439 | -7.3952
3.9556 | -34.0443 | -7.4059
3.9617 | -34.0446 | -7.4164
3.9678 | -34.0449 | -7.427
3.9739 | -34.0452 | -7.4378
3.98 | -34.0455 | -7.4488
3.9861 | -34.0458 | -7.4597
3.9922 | -34.046 | -7.4706
3.9983 | -34.0463 | -7.4811
4.0044 | -34.0466 | -7.4915
4.0105 | -34.0469 | -7.5024
4.0166 | -34.0473 | -7.5133
4.0227 | -34.0476 | -7.5241
4.0288 | -34.0478 | -7.5354
4.0349 | -34.0481 | -7.5461
4.041 | -34.0484 | -7.5569
4.0471 | -34.0486 | -7.5678
4.0532 | -34.0489 | -7.5783
4.0593 | -34.0492 | -7.5893
4.0654 | -34.0495 | -7.6
4.0715 | -34.0498 | -7.6113
4.0776 | -34.0501 | -7.622
4.0838 | -34.0504 | -7.6333
4.0899 | -34.0507 | -7.6439
4.096 | -34.051 | -7.6547
4.1021 | -34.0513 | -7.6653
4.1082 | -34.0516 | -7.6759
4.1143 | -34.0518 | -7.6864
4.1204 | -34.0521 | -7.6969
4.1265 | -34.0524 | -7.7074
4.1326 | -34.0527 | -7.718
4.1387 | -34.053 | -7.7287
4.1448 | -34.0533 | -7.7395
4.1509 | -34.0535 | -7.7503
4.157 | -34.0538 | -7.7615
4.1631 | -34.0541 | -7.7724
4.1692 | -34.0544 | -7.7831
4.1753 | -34.0546 | -7.7938
4.1814 | -34.0549 | -7.8042
4.1875 | -34.0552 | -7.8152
4.1936 | -34.0554 | -7.8261
4.1997 | -34.0557 | -7.8364
4.2058 | -34.0559 | -7.8474
4.2119 | -34.0562 | -7.858
4.218 | -34.0564 | -7.8685
4.2241 | -34.0567 | -7.879
4.2303 | -34.057 | -7.8895
4.2364 | -34.0572 | -7.9002
4.2425 | -34.0574 | -7.9109
4.2486 | -34.0576 | -7.9218
4.2547 | -34.0578 | -7.9322
4.2608 | -34.058 | -7.943
4.2669 | -34.0582 | -7.9534
4.273 | -34.0585 | -7.9641
4.2791 | -34.0587 | -7.9748
4.2852 | -34.059 | -7.9856
4.2913 | -34.0591 | -7.996
4.2974 | -34.0594 | -8.0069
4.3035 | -34.0597 | -8.0179
4.3096 | -34.0599 | -8.0287
4.3157 | -34.0601 | -8.0391
4.3218 | -34.0603 | -8.0496
4.3279 | -34.0606 | -8.0603
4.334 | -34.0608 | -8.0708
4.3401 | -34.061 | -8.081
4.3462 | -34.0613 | -8.0918
4.3523 | -34.0615 | -8.102
4.3584 | -34.0617 | -8.1122
4.3645 | -34.0619 | -8.1227
4.3707 | -34.0622 | -8.133
4.3768 | -34.0624 | -8.1435
4.3829 | -34.0626 | -8.1543
4.389 | -34.0629 | -8.1649
4.3951 | -34.0631 | -8.1752
4.4012 | -34.0633 | -8.186
4.4073 | -34.0635 | -8.1964
4.4134 | -34.0637 | -8.2068
4.4195 | -34.064 | -8.2173
4.4256 | -34.0643 | -8.2276
4.4317 | -34.0645 | -8.238
4.4378 | -34.0647 | -8.2482
4.4439 | -34.0649 | -8.2584
4.45 | -34.0651 | -8.2688
4.4561 | -34.0652 | -8.2793
4.4622 | -34.0654 | -8.2897
4.4683 | -34.0657 | -8.3002
4.4744 | -34.0659 | -8.3109
4.4805 | -34.0661 | -8.321
4.4866 | -34.0664 | -8.3314
4.4927 | -34.0666 | -8.3421
4.4988 | -34.0668 | -8.3524
4.5049 | -34.0671 | -8.3625
4.511 | -34.0673 | -8.3726
4.5172 | -34.0676 | -8.383
4.5233 | -34.0678 | -8.3938
4.5294 | -34.068 | -8.4041
4.5355 | -34.0683 | -8.4145
4.5416 | -34.0685 | -8.4249
4.5477 | -34.0687 | -8.4353
4.5538 | -34.0689 | -8.4453
4.5599 | -34.0692 | -8.4556
4.566 | -34.0694 | -8.4661
4.5721 | -34.0697 | -8.4764
4.5782 | -34.07 | -8.4871
4.5843 | -34.0702 | -8.4975
4.5904 | -34.0705 | -8.508
4.5965 | -34.0707 | -8.5182
4.6026 | -34.071 | -8.529
4.6087 | -34.0712 | -8.5393
4.6148 | -34.0714 | -8.55
4.6209 | -34.0717 | -8.5603
4.627 | -34.072 | -8.5709
4.6331 | -34.0721 | -8.5812
4.6392 | -34.0724 | -8.5916
4.6453 | -34.0726 | -8.602
4.6514 | -34.0729 | -8.6125
4.6576 | -34.0731 | -8.6232
4.6637 | -34.0733 | -8.6333
4.6698 | -34.0735 | -8.6435
4.6759 | -34.0738 | -8.6543
4.682 | -34.074 | -8.6647
4.6881 | -34.0742 | -8.675
4.6942 | -34.0745 | -8.6854
4.7003 | -34.0747 | -8.6959
4.7064 | -34.0749 | -8.7061
4.7125 | -34.0752 | -8.7168
4.7186 | -34.0754 | -8.727
4.7247 | -34.0757 | -8.7375
4.7308 | -34.076 | -8.7479
4.7369 | -34.0762 | -8.7584
4.743 | -34.0765 | -8.7693
4.7491 | -34.0767 | -8.78
4.7552 | -34.077 | -8.7906
4.7613 | -34.0772 | -8.8011
4.7674 | -34.0774 | -8.8114
4.7735 | -34.0776 | -8.8223
4.7796 | -34.078 | -8.833
4.7857 | -34.0782 | -8.8436
4.7918 | -34.0785 | -8.8538
4.7979 | -34.0788 | -8.8642
4.8041 | -34.079 | -8.8747
4.8102 | -34.0793 | -8.8849
4.8163 | -34.0796 | -8.8953
4.8224 | -34.0798 | -8.9057
4.8285 | -34.0801 | -8.9159
4.8346 | -34.0804 | -8.9262
4.8407 | -34.0806 | -8.9368
4.8468 | -34.0809 | -8.9475
4.8529 | -34.0812 | -8.9582
4.859 | -34.0814 | -8.969
4.8651 | -34.0817 | -8.9793
4.8712 | -34.0819 | -8.9896
4.8773 | -34.0822 | -9.0002
4.8834 | -34.0825 | -9.0107
4.8895 | -34.0827 | -9.0209
4.8956 | -34.0829 | -9.0314
4.9017 | -34.0831 | -9.0418
4.9078 | -34.0834 | -9.0526
4.9139 | -34.0837 | -9.0627
4.92 | -34.084 | -9.0729
4.9261 | -34.0843 | -9.0831
4.9322 | -34.0846 | -9.0935
4.9383 | -34.0849 | -9.104
4.9445 | -34.0852 | -9.1144
4.9506 | -34.0854 | -9.1247
4.9567 | -34.0857 | -9.1351
4.9628 | -34.086 | -9.1455
4.9689 | -34.0862 | -9.1564
4.975 | -34.0865 | -9.1671
4.9811 | -34.0868 | -9.1773
4.9872 | -34.0871 | -9.1877
4.9933 | -34.0874 | -9.1984
4.9994 | -34.0877 | -9.2093
5.0055 | -34.088 | -9.2198
5.0116 | -34.0883 | -9.2307
5.0177 | -34.0887 | -9.2415
5.0238 | -34.089 | -9.2523
5.0299 | -34.0893 | -9.2632
5.036 | -34.0896 | -9.2737
5.0421 | -34.0899 | -9.2841
5.0482 | -34.0902 | -9.2946
5.0543 | -34.0905 | -9.3049
5.0604 | -34.0908 | -9.3154
5.0665 | -34.0911 | -9.3257
5.0726 | -34.0914 | -9.3365
5.0787 | -34.0917 | -9.3473
5.0848 | -34.0921 | -9.3579
5.091 | -34.0923 | -9.3683
5.0971 | -34.0927 | -9.3785
5.1032 | -34.093 | -9.3894
5.1093 | -34.0933 | -9.3999
5.1154 | -34.0936 | -9.4102
5.1215 | -34.0938 | -9.4206
5.1276 | -34.0941 | -9.4309
5.1337 | -34.0945 | -9.4412
5.1398 | -34.0948 | -9.4515
5.1459 | -34.0951 | -9.4623
5.152 | -34.0954 | -9.473
5.1581 | -34.0957 | -9.4837
5.1642 | -34.0961 | -9.4943
5.1703 | -34.0963 | -9.5048
5.1764 | -34.0967 | -9.5152
5.1825 | -34.097 | -9.5255
5.1886 | -34.0973 | -9.5357
5.1947 | -34.0976 | -9.5462
5.2008 | -34.0979 | -9.5568
5.2069 | -34.0982 | -9.5675
5.213 | -34.0986 | -9.5778
5.2191 | -34.0989 | -9.5885
5.2252 | -34.0992 | -9.599
5.2314 | -34.0995 | -9.6098
5.2375 | -34.0998 | -9.62
5.2436 | -34.1001 | -9.6306
5.2497 | -34.1004 | -9.6408
5.2558 | -34.1008 | -9.6517
5.2619 | -34.1011 | -9.6625
5.268 | -34.1014 | -9.6732
5.2741 | -34.1018 | -9.6839
5.2802 | -34.1021 | -9.6947
5.2863 | -34.1024 | -9.7052
5.2924 | -34.1028 | -9.716
5.2985 | -34.1031 | -9.7266
5.3046 | -34.1034 | -9.7372
5.3107 | -34.1037 | -9.7479
5.3168 | -34.1041 | -9.7585
5.3229 | -34.1044 | -9.7691
5.329 | -34.1047 | -9.7797
5.3351 | -34.105 | -9.79
5.3412 | -34.1053 | -9.8005
5.3473 | -34.1057 | -9.8111
5.3534 | -34.106 | -9.8218
5.3595 | -34.1063 | -9.8324
5.3656 | -34.1067 | -9.8429
5.3717 | -34.107 | -9.8538
5.3779 | -34.1073 | -9.8647
5.384 | -34.1076 | -9.8753
5.3901 | -34.108 | -9.8861
5.3962 | -34.1083 | -9.8965
5.4023 | -34.1086 | -9.9073
5.4084 | -34.109 | -9.9178
5.4145 | -34.1093 | -9.9282
5.4206 | -34.1096 | -9.9388
5.4267 | -34.11 | -9.9497
5.4328 | -34.1103 | -9.9603
5.4389 | -34.1106 | -9.9706
5.445 | -34.1109 | -9.9816
5.4511 | -34.1113 | -9.992
5.4572 | -34.1116 | -10.0027
5.4633 | -34.112 | -10.0133
5.4694 | -34.1122 | -10.0239
5.4755 | -34.1126 | -10.0342
5.4816 | -34.1129 | -10.0449
5.4877 | -34.1132 | -10.0553
5.4938 | -34.1135 | -10.0657
5.4999 | -34.1139 | -10.0761
5.506 | -34.1143 | -10.0867
5.5121 | -34.1146 | -10.0971
5.5183 | -34.1149 | -10.1074
5.5244 | -34.1152 | -10.1182
5.5305 | -34.1156 | -10.1287
5.5366 | -34.1159 | -10.139
5.5427 | -34.1162 | -10.1497
5.5488 | -34.1165 | -10.1603
5.5549 | -34.1168 | -10.1709
5.561 | -34.1171 | -10.1817
5.5671 | -34.1175 | -10.1921
5.5732 | -34.1178 | -10.2029
5.5793 | -34.1182 | -10.2137
5.5854 | -34.1185 | -10.2242
5.5915 | -34.1188 | -10.2351
5.5976 | -34.1191 | -10.2458
5.6037 | -34.1194 | -10.2564
5.6098 | -34.1197 | -10.2669
5.6159 | -34.1201 | -10.2773
5.622 | -34.1204 | -10.2879
5.6281 | -34.1207 | -10.2985
5.6342 | -34.1211 | -10.3094
5.6403 | -34.1214 | -10.3201
5.6464 | -34.1217 | -10.3308
5.6525 | -34.122 | -10.3415
5.6586 | -34.1224 | -10.3517
5.6648 | -34.1227 | -10.3624
5.6709 | -34.1231 | -10.3731
5.677 | -34.1234 | -10.3837
5.6831 | -34.1237 | -10.3944
5.6892 | -34.124 | -10.4051
5.6953 | -34.1243 | -10.4157
5.7014 | -34.1247 | -10.4263
5.7075 | -34.125 | -10.4369
5.7136 | -34.1253 | -10.4474
5.7197 | -34.1256 | -10.458
5.7258 | -34.1259 | -10.4683
5.7319 | -34.1262 | -10.4789
5.738 | -34.1266 | -10.4895
5.7441 | -34.1269 | -10.5
5.7502 | -34.1273 | -10.5104
5.7563 | -34.1276 | -10.5209
5.7624 | -34.1279 | -10.5317
5.7685 | -34.1283 | -10.5423
5.7746 | -34.1286 | -10.5531
5.7807 | -34.1288 | -10.5637
5.7868 | -34.1291 | -10.5745
5.7929 | -34.1294 | -10.585
5.799 | -34.1298 | -10.5954
5.8052 | -34.1301 | -10.6058
5.8113 | -34.1304 | -10.6162
5.8174 | -34.1307 | -10.6267
5.8235 | -34.131 | -10.6371
5.8296 | -34.1313 | -10.6479
5.8357 | -34.1316 | -10.6583
5.8418 | -34.132 | -10.669
5.8479 | -34.1322 | -10.6796
5.854 | -34.1326 | -10.6903
5.8601 | -34.1329 | -10.7008
5.8662 | -34.1332 | -10.7115
5.8723 | -34.1335 | -10.7219
5.8784 | -34.1338 | -10.7323
5.8845 | -34.1341 | -10.7427
5.8906 | -34.1344 | -10.7534
5.8967 | -34.1348 | -10.7637
5.9028 | -34.1351 | -10.7741
5.9089 | -34.1355 | -10.7845
5.915 | -34.1358 | -10.7948
5.9211 | -34.1361 | -10.8053
5.9272 | -34.1364 | -10.8155
5.9333 | -34.1367 | -10.8259
5.9394 | -34.1371 | -10.8364
5.9455 | -34.1374 | -10.8467
5.9517 | -34.1378 | -10.857
5.9578 | -34.1382 | -10.8676
5.9639 | -34.1385 | -10.8778
5.97 | -34.1388 | -10.8884
5.9761 | -34.1392 | -10.8987
5.9822 | -34.1395 | -10.9094
5.9883 | -34.1398 | -10.9199
5.9944 | -34.1402 | -10.9305
6.0005 | -34.1405 | -10.9411
6.0066 | -34.1408 | -10.9514
6.0127 | -34.1412 | -10.9616
6.0188 | -34.1415 | -10.9723
6.0249 | -34.1418 | -10.9828
6.031 | -34.1421 | -10.9931
6.0371 | -34.1425 | -11.0036
6.0432 | -34.1428 | -11.0139
6.0493 | -34.1431 | -11.0243
6.0554 | -34.1435 | -11.0346
6.0615 | -34.1438 | -11.045
6.0676 | -34.1442 | -11.0557
6.0737 | -34.1445 | -11.0662
6.0798 | -34.1448 | -11.0767
6.0859 | -34.1452 | -11.0871
6.0921 | -34.1455 | -11.0973
6.0982 | -34.1458 | -11.1077
6.1043 | -34.1462 | -11.118
6.1104 | -34.1465 | -11.1287
6.1165 | -34.1469 | -11.1387
6.1226 | -34.1472 | -11.1491
6.1287 | -34.1475 | -11.1594
6.1348 | -34.1479 | -11.1698
6.1409 | -34.1482 | -11.1802
6.147 | -34.1485 | -11.1906
6.1531 | -34.1489 | -11.201
6.1592 | -34.1492 | -11.2112
6.1653 | -34.1495 | -11.2215
6.1714 | -34.1498 | -11.232
6.1775 | -34.1502 | -11.2425
6.1836 | -34.1505 | -11.2529
6.1897 | -34.1509 | -11.2633
6.1958 | -34.1512 | -11.274
6.2019 | -34.1516 | -11.2844
6.208 | -34.1519 | -11.2948
6.2141 | -34.1522 | -11.3052
6.2202 | -34.1526 | -11.3157
6.2263 | -34.153 | -11.326
6.2325 | -34.1533 | -11.3362
6.2386 | -34.1536 | -11.3463
6.2447 | -34.154 | -11.3569
6.2508 | -34.1543 | -11.3673
6.2569 | -34.1547 | -11.3776
6.263 | -34.155 | -11.388
6.2691 | -34.1553 | -11.3983
6.2752 | -34.1556 | -11.4088
6.2813 | -34.1559 | -11.4191
6.2874 | -34.1563 | -11.4294
6.2935 | -34.1566 | -11.4397
6.2996 | -34.1569 | -11.45
6.3057 | -34.1573 | -11.4605
6.3118 | -34.1576 | -11.4709
6.3179 | -34.1579 | -11.4813
6.324 | -34.1582 | -11.4917
6.3301 | -34.1586 | -11.5021
6.3362 | -34.1589 | -11.5123
6.3423 | -34.1593 | -11.5229
6.3484 | -34.1596 | -11.5333
6.3545 | -34.1599 | -11.5434
6.3606 | -34.1603 | -11.5535
6.3667 | -34.1606 | -11.5638
6.3728 | -34.1609 | -11.5741
6.379 | -34.1612 | -11.5845
6.3851 | -34.1616 | -11.5946
6.3912 | -34.1619 | -11.6049
6.3973 | -34.1622 | -11.6153
6.4034 | -34.1625 | -11.6255
6.4095 | -34.1629 | -11.6358
6.4156 | -34.1632 | -11.6462
6.4217 | -34.1635 | -11.6565
6.4278 | -34.1639 | -11.6669
6.4339 | -34.1642 | -11.6771
6.44 | -34.1646 | -11.6873
6.4461 | -34.1649 | -11.6978
6.4522 | -34.1652 | -11.708
6.4583 | -34.1655 | -11.7184
6.4644 | -34.1658 | -11.7288
6.4705 | -34.1662 | -11.7389
6.4766 | -34.1665 | -11.7492
6.4827 | -34.1668 | -11.7596
6.4888 | -34.1671 | -11.7697
6.4949 | -34.1675 | -11.7798
6.501 | -34.1678 | -11.7902
6.5071 | -34.1681 | -11.8003
6.5132 | -34.1685 | -11.8107
6.5194 | -34.1688 | -11.821
6.5255 | -34.1692 | -11.8311
6.5316 | -34.1694 | -11.8416
6.5377 | -34.1698 | -11.8519
6.5438 | -34.1701 | -11.862
6.5499 | -34.1704 | -11.8722
6.556 | -34.1708 | -11.8825
6.5621 | -34.1711 | -11.893
6.5682 | -34.1714 | -11.9033
6.5743 | -34.1718 | -11.9136
6.5804 | -34.1721 | -11.924
6.5865 | -34.1724 | -11.9343
6.5926 | -34.1728 | -11.9445
6.5987 | -34.1731 | -11.9546
6.6048 | -34.1734 | -11.9649
6.6109 | -34.1738 | -11.9753
6.617 | -34.1741 | -11.9857
6.6231 | -34.1744 | -11.996
6.6292 | -34.1748 | -12.0063
6.6353 | -34.1751 | -12.0169
6.6414 | -34.1754 | -12.0271
6.6475 | -34.1758 | -12.0374
6.6536 | -34.1761 | -12.0476
6.6597 | -34.1764 | -12.0577
6.6659 | -34.1767 | -12.068
6.672 | -34.1771 | -12.0782
6.6781 | -34.1774 | -12.0885
6.6842 | -34.1778 | -12.0986
6.6903 | -34.1781 | -12.1088
6.6964 | -34.1785 | -12.1191
6.7025 | -34.1788 | -12.1292
6.7086 | -34.1791 | -12.1394
6.7147 | -34.1794 | -12.1498
6.7208 | -34.1798 | -12.1601
6.7269 | -34.1801 | -12.1703
6.733 | -34.1804 | -12.1804
6.7391 | -34.1808 | -12.1907
6.7452 | -34.1811 | -12.2007
6.7513 | -34.1814 | -12.2108
6.7574 | -34.1818 | -12.2211
6.7635 | -34.1821 | -12.2311
6.7696 | -34.1825 | -12.2413
6.7757 | -34.1828 | -12.2515
6.7818 | -34.1832 | -12.2618
6.7879 | -34.1835 | -12.2719
6.794 | -34.1839 | -12.282
6.8001 | -34.1842 | -12.2922
6.8063 | -34.1845 | -12.3024
6.8124 | -34.1848 | -12.3128
6.8185 | -34.1851 | -12.3233
6.8246 | -34.1854 | -12.3334
6.8307 | -34.1858 | -12.3435
6.8368 | -34.1861 | -12.354
6.8429 | -34.1865 | -12.3643
6.849 | -34.1868 | -12.3745
6.8551 | -34.1872 | -12.3847
6.8612 | -34.1875 | -12.3949
6.8673 | -34.1878 | -12.4051
6.8734 | -34.1882 | -12.4155
6.8795 | -34.1885 | -12.4257
6.8856 | -34.1888 | -12.4359
6.8917 | -34.1892 | -12.4461
6.8978 | -34.1895 | -12.4561
6.9039 | -34.1898 | -12.4661
6.91 | -34.1902 | -12.4763
6.9161 | -34.1906 | -12.4864
6.9222 | -34.1909 | -12.4966
6.9283 | -34.1913 | -12.5065
6.9344 | -34.1916 | -12.5167
6.9405 | -34.192 | -12.5268
6.9466 | -34.1923 | -12.5372
6.9528 | -34.1927 | -12.5476
6.9589 | -34.1931 | -12.5579
6.965 | -34.1934 | -12.5682
6.9711 | -34.1938 | -12.5785
6.9772 | -34.1941 | -12.5887
6.9833 | -34.1945 | -12.5988
6.9894 | -34.1948 | -12.6093
6.9955 | -34.1952 | -12.6193
7.0016 | -34.1955 | -12.6294
7.0077 | -34.1959 | -12.6396
7.0138 | -34.1962 | -12.6497
7.0199 | -34.1965 | -12.6598
7.026 | -34.1969 | -12.67
7.0321 | -34.1973 | -12.6799
7.0382 | -34.1976 | -12.6899
7.0443 | -34.198 | -12.7002
7.0504 | -34.1983 | -12.7104
7.0565 | -34.1987 | -12.7206
7.0626 | -34.199 | -12.7305
7.0687 | -34.1994 | -12.7403
7.0748 | -34.1997 | -12.7502
7.0809 | -34.2001 | -12.7603
7.087 | -34.2004 | -12.7706
7.0932 | -34.2008 | -12.7808
7.0993 | -34.2011 | -12.7909
7.1054 | -34.2015 | -12.8009
7.1115 | -34.2018 | -12.8108
7.1176 | -34.2022 | -12.8209
7.1237 | -34.2025 | -12.831
7.1298 | -34.2029 | -12.841
7.1359 | -34.2032 | -12.8512
7.142 | -34.2035 | -12.861
7.1481 | -34.2039 | -12.871
7.1542 | -34.2042 | -12.8811
7.1603 | -34.2046 | -12.8913
7.1664 | -34.205 | -12.9013
7.1725 | -34.2053 | -12.9114
7.1786 | -34.2057 | -12.9215
7.1847 | -34.206 | -12.9317
7.1908 | -34.2064 | -12.9419
7.1969 | -34.2067 | -12.9522
7.203 | -34.2071 | -12.962
7.2091 | -34.2074 | -12.972
7.2152 | -34.2078 | -12.9821
7.2213 | -34.2082 | -12.992
7.2274 | -34.2085 | -13.0021
7.2335 | -34.2089 | -13.012
7.2397 | -34.2092 | -13.022
7.2458 | -34.2096 | -13.0319
7.2519 | -34.21 | -13.0419
7.258 | -34.2103 | -13.0522
7.2641 | -34.2107 | -13.0623
7.2702 | -34.211 | -13.0725
7.2763 | -34.2114 | -13.0826
7.2824 | -34.2118 | -13.0927
7.2885 | -34.2121 | -13.1028
7.2946 | -34.2125 | -13.1128
7.3007 | -34.2128 | -13.123
7.3068 | -34.2131 | -13.1332
7.3129 | -34.2135 | -13.1433
7.319 | -34.2139 | -13.1534
7.3251 | -34.2142 | -13.1632
7.3312 | -34.2146 | -13.1733
7.3373 | -34.215 | -13.1832
7.3434 | -34.2153 | -13.1931
7.3495 | -34.2157 | -13.203
7.3556 | -34.2161 | -13.2131
7.3617 | -34.2165 | -13.2234
7.3678 | -34.2168 | -13.2333
7.3739 | -34.2172 | -13.2435
7.3801 | -34.2175 | -13.2538
7.3862 | -34.2179 | -13.2638
7.3923 | -34.2183 | -13.274
7.3984 | -34.2186 | -13.2838
7.4045 | -34.219 | -13.2938
7.4106 | -34.2194 | -13.3038
7.4167 | -34.2197 | -13.3139
7.4228 | -34.2201 | -13.324
7.4289 | -34.2205 | -13.334
7.435 | -34.2208 | -13.3438
7.4411 | -34.2212 | -13.3539
7.4472 | -34.2216 | -13.364
7.4533 | -34.2219 | -13.3737
7.4594 | -34.2223 | -13.384
7.4655 | -34.2227 | -13.3939
7.4716 | -34.223 | -13.4041
7.4777 | -34.2234 | -13.4141
7.4838 | -34.2238 | -13.4239
7.4899 | -34.2241 | -13.4338
7.496 | -34.2245 | -13.4436
7.5021 | -34.2248 | -13.4536
7.5082 | -34.2252 | -13.4636
7.5143 | -34.2256 | -13.4737
7.5204 | -34.2259 | -13.4838
7.5266 | -34.2263 | -13.4936
7.5327 | -34.2266 | -13.5035
7.5388 | -34.227 | -13.5138
7.5449 | -34.2274 | -13.524
7.551 | -34.2278 | -13.5339
7.5571 | -34.2281 | -13.544
7.5632 | -34.2285 | -13.554
7.5693 | -34.2289 | -13.5639
7.5754 | -34.2292 | -13.5737
7.5815 | -34.2296 | -13.5838
7.5876 | -34.23 | -13.594
7.5937 | -34.2304 | -13.6041
7.5998 | -34.2308 | -13.614
7.6059 | -34.2311 | -13.6243
7.612 | -34.2315 | -13.6342
7.6181 | -34.2319 | -13.6441
7.6242 | -34.2322 | -13.6544
7.6303 | -34.2326 | -13.6646
7.6364 | -34.2329 | -13.6747
7.6425 | -34.2333 | -13.6847
7.6486 | -34.2337 | -13.6948
7.6547 | -34.2341 | -13.7047
7.6608 | -34.2344 | -13.7147
7.667 | -34.2348 | -13.7247
7.6731 | -34.2351 | -13.7348
7.6792 | -34.2355 | -13.745
7.6853 | -34.2359 | -13.7551
7.6914 | -34.2362 | -13.7653
7.6975 | -34.2366 | -13.7753
7.7036 | -34.237 | -13.7854
7.7097 | -34.2374 | -13.7955
7.7158 | -34.2378 | -13.8056
7.7219 | -34.2381 | -13.8155
7.728 | -34.2385 | -13.8255
7.7341 | -34.2389 | -13.8356
7.7402 | -34.2393 | -13.8457
7.7463 | -34.2396 | -13.8558
7.7524 | -34.24 | -13.8661
7.7585 | -34.2404 | -13.8761
7.7646 | -34.2407 | -13.8861
7.7707 | -34.2411 | -13.8963
7.7768 | -34.2414 | -13.9064
7.7829 | -34.2418 | -13.9167
7.789 | -34.2422 | -13.927
7.7951 | -34.2425 | -13.937
7.8012 | -34.2429 | -13.9471
7.8073 | -34.2433 | -13.9573
7.8135 | -34.2437 | -13.9673
7.8196 | -34.2441 | -13.9775
7.8257 | -34.2444 | -13.9878
7.8318 | -34.2448 | -13.9979
7.8379 | -34.2452 | -14.0079
7.844 | -34.2455 | -14.0179
7.8501 | -34.2459 | -14.0281
7.8562 | -34.2463 | -14.0384
7.8623 | -34.2467 | -14.0484
7.8684 | -34.247 | -14.0585
7.8745 | -34.2474 | -14.0687
7.8806 | -34.2477 | -14.0789
7.8867 | -34.2481 | -14.0887
7.8928 | -34.2485 | -14.0986
7.8989 | -34.2489 | -14.1087
7.905 | -34.2492 | -14.1189
7.9111 | -34.2496 | -14.1289
7.9172 | -34.25 | -14.139
7.9233 | -34.2503 | -14.1492
7.9294 | -34.2507 | -14.1594
7.9355 | -34.2511 | -14.1692
7.9416 | -34.2514 | -14.1794
7.9477 | -34.2518 | -14.1893
7.9539 | -34.2521 | -14.1993
7.96 | -34.2525 | -14.2094
7.9661 | -34.2529 | -14.2195
7.9722 | -34.2533 | -14.2294
7.9783 | -34.2536 | -14.2397
7.9844 | -34.254 | -14.2497
7.9905 | -34.2544 | -14.26
7.9966 | -34.2548 | -14.2701
8.0027 | -34.2551 | -14.2802
8.0088 | -34.2555 | -14.2903
8.0149 | -34.2558 | -14.3004
8.021 | -34.2562 | -14.3105
8.0271 | -34.2566 | -14.3208
8.0332 | -34.2569 | -14.3308
8.0393 | -34.2573 | -14.341
8.0454 | -34.2577 | -14.3513
8.0515 | -34.2581 | -14.3614
8.0576 | -34.2585 | -14.3717
8.0637 | -34.2588 | -14.3816
8.0698 | -34.2592 | -14.3916
8.0759 | -34.2595 | -14.4016
8.082 | -34.2599 | -14.4116
8.0881 | -34.2603 | -14.4217
8.0942 | -34.2607 | -14.4319
8.1004 | -34.2611 | -14.442
8.1065 | -34.2614 | -14.452
8.1126 | -34.2618 | -14.4622
8.1187 | -34.2622 | -14.4723
8.1248 | -34.2625 | -14.4822
8.1309 | -34.2629 | -14.4922
8.137 | -34.2633 | -14.502
8.1431 | -34.2636 | -14.5124
8.1492 | -34.264 | -14.5224
8.1553 | -34.2644 | -14.5326
8.1614 | -34.2648 | -14.5428
8.1675 | -34.2651 | -14.5529
8.1736 | -34.2655 | -14.5631
8.1797 | -34.2658 | -14.5733
8.1858 | -34.2662 | -14.5834
8.1919 | -34.2666 | -14.5935
8.198 | -34.2669 | -14.6034
8.2041 | -34.2673 | -14.6136
8.2102 | -34.2677 | -14.6237
8.2163 | -34.2681 | -14.6336
8.2224 | -34.2684 | -14.6436
8.2285 | -34.2688 | -14.6535
8.2346 | -34.2691 | -14.6635
8.2408 | -34.2695 | -14.6735
8.2469 | -34.2699 | -14.6836
8.253 | -34.2703 | -14.6936
8.2591 | -34.2707 | -14.7037
8.2652 | -34.2711 | -14.7139
8.2713 | -34.2714 | -14.724
8.2774 | -34.2718 | -14.7342
8.2835 | -34.2722 | -14.7445
8.2896 | -34.2726 | -14.7544
8.2957 | -34.2729 | -14.7646
8.3018 | -34.2733 | -14.7746
8.3079 | -34.2737 | -14.7846
8.314 | -34.2741 | -14.7947
8.3201 | -34.2745 | -14.8047
8.3262 | -34.2748 | -14.8148
8.3323 | -34.2752 | -14.8249
8.3384 | -34.2756 | -14.835
8.3445 | -34.2759 | -14.845
8.3506 | -34.2763 | -14.8551
8.3567 | -34.2767 | -14.8652
8.3628 | -34.2771 | -14.8753
8.3689 | -34.2774 | -14.8853
8.375 | -34.2778 | -14.8954
8.3812 | -34.2782 | -14.9053
8.3873 | -34.2786 | -14.9153
8.3934 | -34.2789 | -14.9255
8.3995 | -34.2793 | -14.9354
8.4056 | -34.2796 | -14.9455
8.4117 | -34.28 | -14.9556
8.4178 | -34.2804 | -14.9657
8.4239 | -34.2807 | -14.9756
8.43 | -34.2811 | -14.9857
8.4361 | -34.2814 | -14.9959
8.4422 | -34.2818 | -15.0058
8.4483 | -34.2822 | -15.0158
8.4544 | -34.2826 | -15.0258
8.4605 | -34.283 | -15.0358
8.4666 | -34.2833 | -15.046
8.4727 | -34.2837 | -15.0561
8.4788 | -34.2841 | -15.0661
8.4849 | -34.2844 | -15.0761
8.491 | -34.2848 | -15.0862
8.4971 | -34.2852 | -15.0961
8.5032 | -34.2855 | -15.106
8.5093 | -34.2859 | -15.1159
8.5154 | -34.2863 | -15.126
8.5215 | -34.2867 | -15.1359
8.5277 | -34.2871 | -15.146
8.5338 | -34.2875 | -15.1559
8.5399 | -34.2878 | -15.1659
8.546 | -34.2882 | -15.1758
8.5521 | -34.2886 | -15.1858
8.5582 | -34.2889 | -15.1959
8.5643 | -34.2893 | -15.2057
8.5704 | -34.2897 | -15.2155
8.5765 | -34.2901 | -15.2256
8.5826 | -34.2905 | -15.2353
8.5887 | -34.2909 | -15.2452
8.5948 | -34.2913 | -15.2553
8.6009 | -34.2916 | -15.2651
8.607 | -34.292 | -15.2752
8.6131 | -34.2924 | -15.2852
8.6192 | -34.2928 | -15.2951
8.6253 | -34.2932 | -15.305
8.6314 | -34.2935 | -15.3149
8.6375 | -34.2939 | -15.325
8.6436 | -34.2943 | -15.3349
8.6497 | -34.2947 | -15.3448
8.6558 | -34.295 | -15.3548
8.6619 | -34.2954 | -15.3648
8.6681 | -34.2958 | -15.375
8.6742 | -34.2962 | -15.3849
8.6803 | -34.2965 | -15.3948
8.6864 | -34.2969 | -15.4045
8.6925 | -34.2973 | -15.4143
8.6986 | -34.2977 | -15.4243
8.7047 | -34.2981 | -15.4344
8.7108 | -34.2984 | -15.4444
8.7169 | -34.2988 | -15.4544
8.723 | -34.2992 | -15.4642
8.7291 | -34.2996 | -15.474
8.7352 | -34.3 | -15.4841
8.7413 | -34.3003 | -15.4941
8.7474 | -34.3007 | -15.5039
8.7535 | -34.3011 | -15.5139
8.7596 | -34.3015 | -15.5239
8.7657 | -34.3019 | -15.534
8.7718 | -34.3022 | -15.544
8.7779 | -34.3026 | -15.554
8.784 | -34.303 | -15.564
8.7901 | -34.3034 | -15.574
8.7962 | -34.3038 | -15.5837
8.8023 | -34.3042 | -15.5936
8.8084 | -34.3045 | -15.6035
8.8146 | -34.3049 | -15.6134
8.8207 | -34.3053 | -15.6233
8.8268 | -34.3057 | -15.633
8.8329 | -34.3061 | -15.6429
8.839 | -34.3065 | -15.6529
8.8451 | -34.3068 | -15.6627
8.8512 | -34.3072 | -15.6725
8.8573 | -34.3076 | -15.6825
8.8634 | -34.308 | -15.6922
8.8695 | -34.3084 | -15.7023
8.8756 | -34.3088 | -15.7123
8.8817 | -34.3092 | -15.7221
8.8878 | -34.3095 | -15.7322
8.8939 | -34.3099 | -15.7421
8.9 | -34.3103 | -15.7519
8.9061 | -34.3107 | -15.7618
8.9122 | -34.3111 | -15.7717
8.9183 | -34.3115 | -15.7816
8.9244 | -34.3119 | -15.7917
8.9305 | -34.3123 | -15.8014
8.9366 | -34.3127 | -15.8113
8.9427 | -34.3131 | -15.8213
8.9488 | -34.3135 | -15.8311
8.955 | -34.3139 | -15.841
8.9611 | -34.3143 | -15.8511
8.9672 | -34.3147 | -15.8609
8.9733 | -34.3151 | -15.8706
8.9794 | -34.3155 | -15.8804
8.9855 | -34.3159 | -15.8904
8.9916 | -34.3163 | -15.9005
8.9977 | -34.3167 | -15.9103
9.0038 | -34.3171 | -15.9201
9.0099 | -34.3175 | -15.93
9.016 | -34.3179 | -15.9399
9.0221 | -34.3183 | -15.9496
9.0282 | -34.3187 | -15.9593
9.0343 | -34.3191 | -15.9691
9.0404 | -34.3195 | -15.979
9.0465 | -34.3199 | -15.989
9.0526 | -34.3203 | -15.9988
9.0587 | -34.3208 | -16.0087
9.0648 | -34.3212 | -16.0184
9.0709 | -34.3216 | -16.0283
9.077 | -34.322 | -16.0382
9.0831 | -34.3224 | -16.0482
9.0892 | -34.3228 | -16.0581
9.0953 | -34.3232 | -16.0681
9.1015 | -34.3236 | -16.0779
9.1076 | -34.324 | -16.0877
9.1137 | -34.3244 | -16.0975
9.1198 | -34.3248 | -16.1073
9.1259 | -34.3252 | -16.1171
9.132 | -34.3257 | -16.1272
9.1381 | -34.3261 | -16.1369
9.1442 | -34.3265 | -16.1465
9.1503 | -34.3269 | -16.1564
9.1564 | -34.3273 | -16.1662
9.1625 | -34.3277 | -16.1762
9.1686 | -34.3281 | -16.186
9.1747 | -34.3285 | -16.1958
9.1808 | -34.329 | -16.2056
9.1869 | -34.3294 | -16.2155
9.193 | -34.3298 | -16.2254
9.1991 | -34.3302 | -16.2353
9.2052 | -34.3306 | -16.2451
9.2113 | -34.331 | -16.255
9.2174 | -34.3315 | -16.2648
9.2235 | -34.3319 | -16.2746
9.2296 | -34.3323 | -16.2845
9.2357 | -34.3327 | -16.2945
9.2419 | -34.3332 | -16.3042
9.248 | -34.3336 | -16.3141
9.2541 | -34.334 | -16.324
9.2602 | -34.3345 | -16.3339
9.2663 | -34.3349 | -16.3435
9.2724 | -34.3353 | -16.3533
9.2785 | -34.3357 | -16.3631
9.2846 | -34.3361 | -16.3728
9.2907 | -34.3366 | -16.3826
9.2968 | -34.337 | -16.3926
9.3029 | -34.3374 | -16.4025
9.309 | -34.3378 | -16.4124
9.3151 | -34.3382 | -16.4221
9.3212 | -34.3387 | -16.4318
9.3273 | -34.3391 | -16.4415
9.3334 | -34.3395 | -16.4512
9.3395 | -34.3399 | -16.4609
9.3456 | -34.3403 | -16.4706
9.3517 | -34.3407 | -16.4805
9.3578 | -34.3412 | -16.4903
9.3639 | -34.3416 | -16.5
9.37 | -34.342 | -16.5098
9.3761 | -34.3425 | -16.5198
9.3822 | -34.3429 | -16.5296
9.3884 | -34.3433 | -16.5392
9.3945 | -34.3438 | -16.5491
9.4006 | -34.3442 | -16.5589
9.4067 | -34.3446 | -16.5687
9.4128 | -34.3451 | -16.5785
9.4189 | -34.3455 | -16.5885
9.425 | -34.3459 | -16.5982
9.4311 | -34.3464 | -16.608
9.4372 | -34.3468 | -16.6178
9.4433 | -34.3472 | -16.6275
9.4494 | -34.3477 | -16.6372
9.4555 | -34.3481 | -16.6471
9.4616 | -34.3486 | -16.6568
9.4677 | -34.349 | -16.6666
9.4738 | -34.3494 | -16.6764
9.4799 | -34.3498 | -16.6862
9.486 | -34.3503 | -16.696
9.4921 | -34.3507 | -16.7058
9.4982 | -34.3512 | -16.7154
9.5043 | -34.3516 | -16.7252
9.5104 | -34.352 | -16.7348
9.5165 | -34.3524 | -16.7446
9.5226 | -34.3528 | -16.7545
9.5288 | -34.3533 | -16.7642
9.5349 | -34.3537 | -16.774
9.541 | -34.3542 | -16.7836
9.5471 | -34.3546 | -16.7933
9.5532 | -34.355 | -16.8032
9.5593 | -34.3555 | -16.8131
9.5654 | -34.3559 | -16.8228
9.5715 | -34.3563 | -16.8325
9.5776 | -34.3567 | -16.8424
9.5837 | -34.3571 | -16.8521
9.5898 | -34.3576 | -16.8619
9.5959 | -34.358 | -16.8718
9.602 | -34.3584 | -16.8813
9.6081 | -34.3589 | -16.8911
9.6142 | -34.3593 | -16.9008
9.6203 | -34.3597 | -16.9106
9.6264 | -34.3601 | -16.9202
9.6325 | -34.3605 | -16.9301
9.6386 | -34.361 | -16.9397
9.6447 | -34.3614 | -16.9495
9.6508 | -34.3619 | -16.9592
9.6569 | -34.3623 | -16.969
9.663 | -34.3627 | -16.9788
9.6691 | -34.3632 | -16.9886
9.6753 | -34.3636 | -16.9982
9.6814 | -34.3641 | -17.0078
9.6875 | -34.3645 | -17.0174
9.6936 | -34.3649 | -17.0269
9.6997 | -34.3654 | -17.0368
9.7058 | -34.3658 | -17.0465
9.7119 | -34.3663 | -17.0563
9.718 | -34.3667 | -17.0662
9.7241 | -34.3672 | -17.0758
9.7302 | -34.3676 | -17.0856
9.7363 | -34.368 | -17.0952
9.7424 | -34.3685 | -17.1048
9.7485 | -34.3689 | -17.1146
9.7546 | -34.3693 | -17.1241
9.7607 | -34.3697 | -17.134
9.7668 | -34.3702 | -17.1438
9.7729 | -34.3706 | -17.1535
9.779 | -34.3711 | -17.163
9.7851 | -34.3715 | -17.1728
9.7912 | -34.372 | -17.1824
9.7973 | -34.3725 | -17.1923
9.8034 | -34.3729 | -17.2018
9.8095 | -34.3733 | -17.2117
9.8157 | -34.3738 | -17.2214
9.8218 | -34.3742 | -17.2313
9.8279 | -34.3747 | -17.2408
9.834 | -34.3751 | -17.2506
9.8401 | -34.3755 | -17.2603
9.8462 | -34.376 | -17.2698
9.8523 | -34.3764 | -17.2795
9.8584 | -34.3768 | -17.2892
9.8645 | -34.3773 | -17.299
9.8706 | -34.3777 | -17.3088
9.8767 | -34.3781 | -17.3184
9.8828 | -34.3785 | -17.328
9.8889 | -34.379 | -17.3375
9.895 | -34.3794 | -17.3472
9.9011 | -34.3799 | -17.3569
9.9072 | -34.3803 | -17.3666
9.9133 | -34.3807 | -17.3763
9.9194 | -34.3812 | -17.3858
9.9255 | -34.3816 | -17.3954
9.9316 | -34.382 | -17.4051
9.9377 | -34.3825 | -17.4147
9.9438 | -34.3829 | -17.4245
9.9499 | -34.3833 | -17.4343
9.956 | -34.3838 | -17.444
9.9622 | -34.3842 | -17.4538
9.9683 | -34.3846 | -17.4635
9.9744 | -34.385 | -17.4733
9.9805 | -34.3855 | -17.4829
9.9866 | -34.3859 | -17.4925
9.9927 | -34.3864 | -17.5021
9.9988 | -34.3868 | -17.5119
10.0049 | -34.3872 | -17.5215
10.011 | -34.3877 | -17.5313
10.0171 | -34.3881 | -17.5408
10.0232 | -34.3886 | -17.5504
10.0293 | -34.389 | -17.5601
10.0354 | -34.3895 | -17.5698
10.0415 | -34.3899 | -17.5794
10.0476 | -34.3903 | -17.5892
10.0537 | -34.3908 | -17.599
10.0598 | -34.3912 | -17.6086
10.0659 | -34.3916 | -17.6183
10.072 | -34.3921 | -17.6278
10.0781 | -34.3925 | -17.6373
10.0842 | -34.393 | -17.647
10.0903 | -34.3934 | -17.6566
10.0964 | -34.3938 | -17.666
10.1026 | -34.3943 | -17.6757
10.1087 | -34.3947 | -17.6853
10.1148 | -34.3951 | -17.695
10.1209 | -34.3955 | -17.7045
10.127 | -34.396 | -17.7141
10.1331 | -34.3964 | -17.7236
10.1392 | -34.3968 | -17.7331
10.1453 | -34.3973 | -17.7426
10.1514 | -34.3977 | -17.7523
10.1575 | -34.3981 | -17.762
10.1636 | -34.3986 | -17.7714
10.1697 | -34.399 | -17.7811
10.1758 | -34.3995 | -17.7906
10.1819 | -34.3999 | -17.8001
10.188 | -34.4003 | -17.8098
10.1941 | -34.4007 | -17.8194
10.2002 | -34.4012 | -17.829
10.2063 | -34.4016 | -17.8386
10.2124 | -34.402 | -17.8483
10.2185 | -34.4025 | -17.858
10.2246 | -34.4029 | -17.8675
10.2307 | -34.4034 | -17.8771
10.2368 | -34.4038 | -17.8864
10.2429 | -34.4042 | -17.896
10.2491 | -34.4046 | -17.9057
10.2552 | -34.4051 | -17.9153
10.2613 | -34.4055 | -17.9249
10.2674 | -34.4059 | -17.9346
10.2735 | -34.4064 | -17.9443
10.2796 | -34.4068 | -17.9539
10.2857 | -34.4073 | -17.9636
10.2918 | -34.4077 | -17.9733
10.2979 | -34.4081 | -17.9829
10.304 | -34.4086 | -17.9923
10.3101 | -34.409 | -18.002
10.3162 | -34.4094 | -18.0117
10.3223 | -34.4099 | -18.0211
10.3284 | -34.4103 | -18.0307
10.3345 | -34.4107 | -18.0402
10.3406 | -34.4112 | -18.0497
10.3467 | -34.4116 | -18.0592
10.3528 | -34.412 | -18.0687
10.3589 | -34.4124 | -18.0783
10.365 | -34.4129 | -18.088
10.3711 | -34.4133 | -18.0976
10.3772 | -34.4137 | -18.1072
10.3833 | -34.4142 | -18.1168
10.3895 | -34.4147 | -18.1264
10.3956 | -34.4151 | -18.136
10.4017 | -34.4155 | -18.1454
10.4078 | -34.4159 | -18.1549
10.4139 | -34.4164 | -18.1645
10.42 | -34.4168 | -18.1742
10.4261 | -34.4173 | -18.1837
10.4322 | -34.4177 | -18.1933
10.4383 | -34.4181 | -18.2028
10.4444 | -34.4185 | -18.2122
10.4505 | -34.4189 | -18.2219
10.4566 | -34.4194 | -18.2313
10.4627 | -34.4198 | -18.2407
10.4688 | -34.4202 | -18.2502
10.4749 | -34.4206 | -18.2598
10.481 | -34.4211 | -18.2693
10.4871 | -34.4215 | -18.2787
10.4932 | -34.422 | -18.2882
10.4993 | -34.4224 | -18.2978
10.5054 | -34.4228 | -18.3074
10.5115 | -34.4233 | -18.3168
10.5176 | -34.4237 | -18.3262
10.5237 | -34.4242 | -18.3359
10.5298 | -34.4246 | -18.3453
10.536 | -34.425 | -18.3549
10.5421 | -34.4255 | -18.3643
10.5482 | -34.4259 | -18.3739
10.5543 | -34.4264 | -18.3833
10.5604 | -34.4268 | -18.3927
10.5665 | -34.4272 | -18.4021
10.5726 | -34.4276 | -18.4116
10.5787 | -34.4281 | -18.4211
10.5848 | -34.4285 | -18.4303
10.5909 | -34.4289 | -18.4399
10.597 | -34.4294 | -18.4495
10.6031 | -34.4298 | -18.4592
10.6092 | -34.4302 | -18.4686
10.6153 | -34.4307 | -18.4782
10.6214 | -34.4311 | -18.4875
10.6275 | -34.4315 | -18.497
10.6336 | -34.432 | -18.5065
10.6397 | -34.4324 | -18.516
10.6458 | -34.4328 | -18.5255
10.6519 | -34.4332 | -18.535
10.658 | -34.4337 | -18.5444
10.6641 | -34.4341 | -18.554
10.6702 | -34.4345 | -18.5636
10.6764 | -34.435 | -18.5731
10.6825 | -34.4354 | -18.5825
10.6886 | -34.4358 | -18.5918
10.6947 | -34.4363 | -18.6013
10.7008 | -34.4367 | -18.6108
10.7069 | -34.4372 | -18.6203
10.713 | -34.4376 | -18.6298
10.7191 | -34.4381 | -18.6393
10.7252 | -34.4385 | -18.6485
10.7313 | -34.4389 | -18.6579
10.7374 | -34.4393 | -18.6674
10.7435 | -34.4398 | -18.6769
10.7496 | -34.4402 | -18.6862
10.7557 | -34.4406 | -18.6957
10.7618 | -34.4411 | -18.7049
10.7679 | -34.4415 | -18.7143
10.774 | -34.4419 | -18.7239
10.7801 | -34.4424 | -18.7334
10.7862 | -34.4428 | -18.7427
10.7923 | -34.4432 | -18.7522
10.7984 | -34.4437 | -18.7615
10.8045 | -34.4441 | -18.7708
10.8106 | -34.4445 | -18.7804
10.8168 | -34.4449 | -18.7897
10.8229 | -34.4454 | -18.7991
10.829 | -34.4458 | -18.8085
10.8351 | -34.4463 | -18.8177
10.8412 | -34.4467 | -18.827
10.8473 | -34.4472 | -18.8363
10.8534 | -34.4476 | -18.8457
10.8595 | -34.448 | -18.8551
10.8656 | -34.4485 | -18.8643
10.8717 | -34.4489 | -18.8737
10.8778 | -34.4494 | -18.8832
10.8839 | -34.4498 | -18.8925
10.89 | -34.4502 | -18.9019
10.8961 | -34.4507 | -18.9113
10.9022 | -34.4511 | -18.9207
10.9083 | -34.4515 | -18.93
10.9144 | -34.452 | -18.9393
10.9205 | -34.4524 | -18.9488
10.9266 | -34.4528 | -18.9581
10.9327 | -34.4533 | -18.9675
10.9388 | -34.4537 | -18.9767
10.9449 | -34.4542 | -18.9861
10.951 | -34.4546 | -18.9955
10.9571 | -34.455 | -19.0049
10.9633 | -34.4555 | -19.0144
10.9694 | -34.4559 | -19.0239
10.9755 | -34.4564 | -19.0333
10.9816 | -34.4568 | -19.0428
10.9877 | -34.4572 | -19.052
10.9938 | -34.4577 | -19.0616
10.9999 | -34.4581 | -19.0708
11.006 | -34.4586 | -19.0799
11.0121 | -34.459 | -19.0893
11.0182 | -34.4594 | -19.0987
11.0243 | -34.4599 | -19.108
11.0304 | -34.4603 | -19.1174
11.0365 | -34.4607 | -19.1266
11.0426 | -34.4612 | -19.1359
11.0487 | -34.4617 | -19.1453
11.0548 | -34.4621 | -19.1546
11.0609 | -34.4626 | -19.164
11.067 | -34.463 | -19.1733
11.0731 | -34.4635 | -19.1826
11.0792 | -34.4639 | -19.1918
11.0853 | -34.4644 | -19.2012
11.0914 | -34.4648 | -19.2106
11.0975 | -34.4653 | -19.2199
11.1037 | -34.4657 | -19.2292
11.1098 | -34.4661 | -19.2387
11.1159 | -34.4666 | -19.2478
11.122 | -34.467 | -19.257
11.1281 | -34.4674 | -19.2664
11.1342 | -34.4679 | -19.2757
11.1403 | -34.4683 | -19.2851
11.1464 | -34.4688 | -19.2945
11.1525 | -34.4692 | -19.3036
11.1586 | -34.4696 | -19.313
11.1647 | -34.4701 | -19.3223
11.1708 | -34.4705 | -19.3315
11.1769 | -34.471 | -19.3408
11.183 | -34.4714 | -19.3499
11.1891 | -34.4719 | -19.3591
11.1952 | -34.4723 | -19.3684
11.2013 | -34.4728 | -19.3778
11.2074 | -34.4732 | -19.3871
11.2135 | -34.4737 | -19.3965
11.2196 | -34.4741 | -19.4056
11.2257 | -34.4746 | -19.415
11.2318 | -34.475 | -19.4242
11.2379 | -34.4755 | -19.4336
11.244 | -34.476 | -19.443
11.2502 | -34.4764 | -19.4522
11.2563 | -34.4769 | -19.4615
11.2624 | -34.4773 | -19.4707
11.2685 | -34.4778 | -19.48
11.2746 | -34.4782 | -19.4892
11.2807 | -34.4787 | -19.4986
11.2868 | -34.4791 | -19.5079
11.2929 | -34.4796 | -19.517
11.299 | -34.4801 | -19.5264
11.3051 | -34.4805 | -19.5358
11.3112 | -34.481 | -19.5452
11.3173 | -34.4814 | -19.5543
11.3234 | -34.4819 | -19.5633
11.3295 | -34.4823 | -19.5728
11.3356 | -34.4828 | -19.5822
11.3417 | -34.4832 | -19.5914
11.3478 | -34.4836 | -19.6008
11.3539 | -34.4841 | -19.6101
11.36 | -34.4845 | -19.6193
11.3661 | -34.485 | -19.6284
11.3722 | -34.4854 | -19.6375
11.3783 | -34.4859 | -19.6469
11.3844 | -34.4863 | -19.6562
11.3906 | -34.4868 | -19.6654
11.3967 | -34.4872 | -19.6745
11.4028 | -34.4877 | -19.6839
11.4089 | -34.4881 | -19.6929
11.415 | -34.4885 | -19.7022
11.4211 | -34.489 | -19.7116
11.4272 | -34.4894 | -19.7209
11.4333 | -34.4898 | -19.7301
11.4394 | -34.4903 | -19.7393
11.4455 | -34.4907 | -19.7485
11.4516 | -34.4911 | -19.7577
11.4577 | -34.4916 | -19.7669
11.4638 | -34.4921 | -19.776
11.4699 | -34.4925 | -19.7853
11.476 | -34.493 | -19.7947
11.4821 | -34.4934 | -19.8038
11.4882 | -34.4939 | -19.8132
11.4943 | -34.4944 | -19.8225
11.5004 | -34.4948 | -19.8315
11.5065 | -34.4953 | -19.8407
11.5126 | -34.4957 | -19.8498
11.5187 | -34.4962 | -19.859
11.5248 | -34.4966 | -19.8684
11.5309 | -34.4971 | -19.8775
11.5371 | -34.4975 | -19.8866
11.5432 | -34.498 | -19.8958
11.5493 | -34.4984 | -19.9051
11.5554 | -34.4989 | -19.9143
11.5615 | -34.4993 | -19.9237
11.5676 | -34.4998 | -19.9329
11.5737 | -34.5003 | -19.9419
11.5798 | -34.5007 | -19.951
11.5859 | -34.5012 | -19.9602
11.592 | -34.5016 | -19.9693
11.5981 | -34.502 | -19.9785
11.6042 | -34.5025 | -19.9877
11.6103 | -34.5029 | -19.9967
11.6164 | -34.5034 | -20.0057
11.6225 | -34.5038 | -20.0149
11.6286 | -34.5043 | -20.024
11.6347 | -34.5047 | -20.0333
11.6408 | -34.5052 | -20.0422
11.6469 | -34.5056 | -20.0514
11.653 | -34.5061 | -20.0606
11.6591 | -34.5065 | -20.0698
11.6652 | -34.507 | -20.0788
11.6713 | -34.5074 | -20.088
11.6775 | -34.5079 | -20.0969
11.6836 | -34.5084 | -20.1061
11.6897 | -34.5088 | -20.1153
11.6958 | -34.5093 | -20.1245
11.7019 | -34.5097 | -20.1338
11.708 | -34.5102 | -20.1429
11.7141 | -34.5107 | -20.1519
11.7202 | -34.5111 | -20.161
11.7263 | -34.5115 | -20.1702
11.7324 | -34.512 | -20.1795
11.7385 | -34.5124 | -20.1887
11.7446 | -34.5129 | -20.1977
11.7507 | -34.5134 | -20.2066
11.7568 | -34.5138 | -20.2158
11.7629 | -34.5142 | -20.225
11.769 | -34.5147 | -20.2341
11.7751 | -34.5151 | -20.2432
11.7812 | -34.5156 | -20.2521
11.7873 | -34.5161 | -20.2612
11.7934 | -34.5166 | -20.2703
11.7995 | -34.517 | -20.2794
11.8056 | -34.5175 | -20.2885
11.8117 | -34.5179 | -20.2976
11.8178 | -34.5184 | -20.3065
11.824 | -34.5189 | -20.3155
11.8301 | -34.5193 | -20.3246
11.8362 | -34.5198 | -20.3336
11.8423 | -34.5202 | -20.3427
11.8484 | -34.5207 | -20.3517
11.8545 | -34.5212 | -20.3606
11.8606 | -34.5216 | -20.3697
11.8667 | -34.5221 | -20.3788
11.8728 | -34.5225 | -20.3878
11.8789 | -34.523 | -20.3969
11.885 | -34.5235 | -20.406
11.8911 | -34.5239 | -20.415
11.8972 | -34.5244 | -20.4239
11.9033 | -34.5248 | -20.433
11.9094 | -34.5253 | -20.4419
11.9155 | -34.5257 | -20.451
11.9216 | -34.5262 | -20.4597
11.9277 | -34.5266 | -20.4687
11.9338 | -34.5271 | -20.4779
11.9399 | -34.5275 | -20.487
11.946 | -34.528 | -20.4959
11.9521 | -34.5285 | -20.5049
11.9582 | -34.529 | -20.5138
11.9644 | -34.5294 | -20.5228
11.9705 | -34.5299 | -20.5318
11.9766 | -34.5303 | -20.5408
11.9827 | -34.5308 | -20.5498
11.9888 | -34.5313 | -20.5589
11.9949 | -34.5318 | -20.5676
12.001 | -34.5322 | -20.5766
12.0071 | -34.5327 | -20.5855
12.0132 | -34.5331 | -20.5945
12.0193 | -34.5336 | -20.6036
12.0254 | -34.534 | -20.6125
12.0315 | -34.5345 | -20.6212
12.0376 | -34.535 | -20.6302
12.0437 | -34.5354 | -20.6391
12.0498 | -34.5359 | -20.6482
12.0559 | -34.5363 | -20.6571
12.062 | -34.5368 | -20.6658
12.0681 | -34.5373 | -20.6748
12.0742 | -34.5378 | -20.6837
12.0803 | -34.5382 | -20.6927
12.0864 | -34.5387 | -20.7018
12.0925 | -34.5391 | -20.7107
12.0986 | -34.5396 | -20.7195
12.1047 | -34.5401 | -20.7285
12.1109 | -34.5406 | -20.7374
12.117 | -34.541 | -20.7464
12.1231 | -34.5415 | -20.7553
12.1292 | -34.5419 | -20.7642
12.1353 | -34.5424 | -20.773
12.1414 | -34.5429 | -20.782
12.1475 | -34.5433 | -20.7911
12.1536 | -34.5438 | -20.8001
12.1597 | -34.5443 | -20.809
12.1658 | -34.5448 | -20.8178
12.1719 | -34.5453 | -20.8266
12.178 | -34.5457 | -20.8355
12.1841 | -34.5462 | -20.8445
12.1902 | -34.5466 | -20.8534
12.1963 | -34.5471 | -20.8625
12.2024 | -34.5476 | -20.8712
12.2085 | -34.5481 | -20.8802
12.2146 | -34.5485 | -20.8892
12.2207 | -34.549 | -20.8982
12.2268 | -34.5495 | -20.9071
12.2329 | -34.55 | -20.9159
12.239 | -34.5505 | -20.9244
12.2451 | -34.5509 | -20.9334
12.2513 | -34.5514 | -20.9425
12.2574 | -34.5519 | -20.9513
12.2635 | -34.5524 | -20.9602
12.2696 | -34.5528 | -20.9691
12.2757 | -34.5533 | -20.9779
12.2818 | -34.5538 | -20.9869
12.2879 | -34.5542 | -20.9958
12.294 | -34.5547 | -21.0047
12.3001 | -34.5552 | -21.0135
12.3062 | -34.5557 | -21.0223
12.3123 | -34.5561 | -21.031
12.3184 | -34.5566 | -21.0397
12.3245 | -34.5571 | -21.0485
12.3306 | -34.5575 | -21.0573
12.3367 | -34.558 | -21.0662
12.3428 | -34.5585 | -21.0751
12.3489 | -34.559 | -21.0837
12.355 | -34.5595 | -21.0926
12.3611 | -34.5599 | -21.1016
12.3672 | -34.5604 | -21.1104
12.3733 | -34.5609 | -21.1192
12.3794 | -34.5614 | -21.1278
12.3855 | -34.5618 | -21.1364
12.3916 | -34.5623 | -21.1451
12.3978 | -34.5628 | -21.1542
12.4039 | -34.5633 | -21.163
12.41 | -34.5637 | -21.1718
12.4161 | -34.5642 | -21.1803
12.4222 | -34.5647 | -21.1892
12.4283 | -34.5652 | -21.198
12.4344 | -34.5657 | -21.2069
12.4405 | -34.5662 | -21.216
12.4466 | -34.5666 | -21.2248
12.4527 | -34.5671 | -21.2335
12.4588 | -34.5676 | -21.2424
12.4649 | -34.568 | -21.2511
12.471 | -34.5685 | -21.26
12.4771 | -34.569 | -21.2686
12.4832 | -34.5695 | -21.2772
12.4893 | -34.57 | -21.2858
12.4954 | -34.5705 | -21.2947
12.5015 | -34.5709 | -21.3032
12.5076 | -34.5714 | -21.3123
12.5137 | -34.5719 | -21.3211
12.5198 | -34.5723 | -21.3297
12.5259 | -34.5727 | -21.3385
12.532 | -34.5732 | -21.3472
12.5382 | -34.5737 | -21.3562
12.5443 | -34.5742 | -21.3649
12.5504 | -34.5746 | -21.3738
12.5565 | -34.5751 | -21.3823
12.5626 | -34.5757 | -21.3912
12.5687 | -34.5761 | -21.4
12.5748 | -34.5766 | -21.4091
12.5809 | -34.5771 | -21.4182
12.587 | -34.5776 | -21.4272
12.5931 | -34.5781 | -21.435
12.5992 | -34.5786 | -21.4437
12.6053 | -34.579 | -21.4527
12.6114 | -34.5795 | -21.4612
12.6175 | -34.58 | -21.4702
12.6236 | -34.5804 | -21.4788
12.6297 | -34.5809 | -21.4873
12.6358 | -34.5814 | -21.4961
12.6419 | -34.5819 | -21.5046
12.648 | -34.5824 | -21.5133
12.6541 | -34.5829 | -21.5219
12.6602 | -34.5834 | -21.5306
12.6663 | -34.5838 | -21.5394
12.6724 | -34.5843 | -21.548
12.6785 | -34.5848 | -21.5567
12.6847 | -34.5852 | -21.5652
12.6908 | -34.5858 | -21.5737
12.6969 | -34.5863 | -21.5823
12.703 | -34.5867 | -21.5912
12.7091 | -34.5872 | -21.5999
12.7152 | -34.5877 | -21.6085
12.7213 | -34.5882 | -21.6171
12.7274 | -34.5887 | -21.6258
12.7335 | -34.5892 | -21.6343
12.7396 | -34.5897 | -21.6428
12.7457 | -34.5902 | -21.6518
12.7518 | -34.5906 | -21.6604
12.7579 | -34.5911 | -21.6691
12.764 | -34.5917 | -21.6783
12.7701 | -34.5922 | -21.687
12.7762 | -34.5927 | -21.6955
12.7823 | -34.5932 | -21.704
12.7884 | -34.5937 | -21.7128
12.7945 | -34.5941 | -21.7217
12.8006 | -34.5947 | -21.7307
12.8067 | -34.5951 | -21.7391
12.8128 | -34.5956 | -21.7476
12.8189 | -34.5961 | -21.7558
12.8251 | -34.5966 | -21.7643
12.8312 | -34.597 | -21.7731
12.8373 | -34.5976 | -21.7822
12.8434 | -34.5981 | -21.791
12.8495 | -34.5985 | -21.7996
12.8556 | -34.599 | -21.808
12.8617 | -34.5995 | -21.8169
12.8678 | -34.6 | -21.8258
12.8739 | -34.6005 | -21.8345
12.88 | -34.601 | -21.8429
12.8861 | -34.6014 | -21.8516
12.8922 | -34.6019 | -21.8602
12.8983 | -34.6024 | -21.8689
12.9044 | -34.6029 | -21.8771
12.9105 | -34.6034 | -21.8859
12.9166 | -34.6039 | -21.8946
12.9227 | -34.6043 | -21.9033
12.9288 | -34.6049 | -21.9118
12.9349 | -34.6053 | -21.9203
12.941 | -34.6059 | -21.9288
12.9471 | -34.6063 | -21.9377
12.9532 | -34.6069 | -21.9462
12.9593 | -34.6073 | -21.9553
12.9654 | -34.6079 | -21.964
12.9716 | -34.6084 | -21.9727
12.9777 | -34.6089 | -21.9817
12.9838 | -34.6094 | -21.99
12.9899 | -34.6099 | -21.9987
12.996 | -34.6104 | -22.0071
13.0021 | -34.6109 | -22.0158
13.0082 | -34.6114 | -22.0245
13.0143 | -34.6119 | -22.0331
13.0204 | -34.6124 | -22.0418
13.0265 | -34.6129 | -22.0507
13.0326 | -34.6134 | -22.0592
13.0387 | -34.6139 | -22.0679
13.0448 | -34.6143 | -22.0763
13.0509 | -34.6149 | -22.085
13.057 | -34.6154 | -22.0935
13.0631 | -34.6159 | -22.1021
13.0692 | -34.6164 | -22.1111
13.0753 | -34.6169 | -22.1201
13.0814 | -34.6174 | -22.1292
13.0875 | -34.6179 | -22.1376
13.0936 | -34.6184 | -22.1467
13.0997 | -34.6189 | -22.155
13.1058 | -34.6194 | -22.1633
13.112 | -34.6199 | -22.1716
13.1181 | -34.6205 | -22.1799
13.1242 | -34.621 | -22.1889
13.1303 | -34.6215 | -22.1974
13.1364 | -34.622 | -22.2064
13.1425 | -34.6225 | -22.2153
13.1486 | -34.623 | -22.2239
13.1547 | -34.6236 | -22.2325
13.1608 | -34.6241 | -22.2411
13.1669 | -34.6246 | -22.2497
13.173 | -34.6251 | -22.2584
13.1791 | -34.6256 | -22.2673
13.1852 | -34.6261 | -22.2763
13.1913 | -34.6266 | -22.2846
13.1974 | -34.6271 | -22.2935
13.2035 | -34.6275 | -22.3021
13.2096 | -34.6281 | -22.3105
13.2157 | -34.6286 | -22.3185
13.2218 | -34.6291 | -22.3273
13.2279 | -34.6296 | -22.3357
13.234 | -34.6301 | -22.3443
13.2401 | -34.6306 | -22.3529
13.2462 | -34.6311 | -22.3616
13.2524 | -34.6316 | -22.3699
13.2585 | -34.6321 | -22.3784
13.2646 | -34.6326 | -22.3871
13.2707 | -34.6331 | -22.3958
13.2768 | -34.6336 | -22.4047
13.2829 | -34.6341 | -22.4136
13.289 | -34.6346 | -22.4222
13.2951 | -34.6352 | -22.4308
13.3012 | -34.6357 | -22.4396
13.3073 | -34.6362 | -22.448
13.3134 | -34.6366 | -22.4565
13.3195 | -34.6371 | -22.465
13.3256 | -34.6376 | -22.4738
13.3317 | -34.6382 | -22.4823
13.3378 | -34.6387 | -22.491
13.3439 | -34.6392 | -22.4995
13.35 | -34.6396 | -22.5081
13.3561 | -34.6401 | -22.5164
13.3622 | -34.6405 | -22.5248
13.3683 | -34.6411 | -22.5332
13.3744 | -34.6416 | -22.5415
13.3805 | -34.6421 | -22.5502
13.3866 | -34.6426 | -22.5588
13.3927 | -34.643 | -22.5671
13.3989 | -34.6435 | -22.5759
13.405 | -34.6439 | -22.5844
13.4111 | -34.6444 | -22.5927
13.4172 | -34.645 | -22.6015
13.4233 | -34.6455 | -22.6098
13.4294 | -34.646 | -22.6185
13.4355 | -34.6466 | -22.6273
13.4416 | -34.6471 | -22.6359
13.4477 | -34.6476 | -22.6447
13.4538 | -34.6481 | -22.6531
13.4599 | -34.6485 | -22.6616
13.466 | -34.649 | -22.6702
13.4721 | -34.6495 | -22.6784
13.4782 | -34.65 | -22.6872
13.4843 | -34.6505 | -22.6957
13.4904 | -34.6511 | -22.7044
13.4965 | -34.6515 | -22.7126
13.5026 | -34.6521 | -22.7211
13.5087 | -34.6526 | -22.7291
13.5148 | -34.6531 | -22.738
13.5209 | -34.6535 | -22.7468
13.527 | -34.654 | -22.755
13.5331 | -34.6546 | -22.7634
13.5393 | -34.6551 | -22.7718
13.5454 | -34.6556 | -22.7804
13.5515 | -34.6561 | -22.7889
13.5576 | -34.6566 | -22.7974
13.5637 | -34.6571 | -22.8058
13.5698 | -34.6576 | -22.8144
13.5759 | -34.6581 | -22.8227
13.582 | -34.6586 | -22.8315
13.5881 | -34.6591 | -22.8398
13.5942 | -34.6596 | -22.8484
13.6003 | -34.6602 | -22.8567
13.6064 | -34.6607 | -22.8652
13.6125 | -34.6612 | -22.8735
13.6186 | -34.6617 | -22.882
13.6247 | -34.6622 | -22.8907
13.6308 | -34.6627 | -22.8992
13.6369 | -34.6633 | -22.9079
13.643 | -34.6637 | -22.9164
13.6491 | -34.6643 | -22.925
13.6552 | -34.6648 | -22.9338
13.6613 | -34.6654 | -22.9422
13.6674 | -34.6658 | -22.9506
13.6735 | -34.6662 | -22.9589
13.6796 | -34.6668 | -22.9674
13.6858 | -34.6672 | -22.9758
13.6919 | -34.6677 | -22.9844
13.698 | -34.6683 | -22.9929
13.7041 | -34.6688 | -23.0014
13.7102 | -34.6692 | -23.0098
13.7163 | -34.6697 | -23.018
13.7224 | -34.6702 | -23.0266
13.7285 | -34.6707 | -23.0353
13.7346 | -34.6712 | -23.0436
13.7407 | -34.6716 | -23.0524
13.7468 | -34.6721 | -23.0606
13.7529 | -34.6727 | -23.0692
13.759 | -34.6732 | -23.0777
13.7651 | -34.6736 | -23.0867
13.7712 | -34.6741 | -23.0959
13.7773 | -34.6747 | -23.1043
13.7834 | -34.6751 | -23.1129
13.7895 | -34.6756 | -23.1222
13.7956 | -34.676 | -23.131
13.8017 | -34.6765 | -23.1392
13.8078 | -34.6771 | -23.1483
13.8139 | -34.6775 | -23.1564
13.82 | -34.678 | -23.1647
13.8262 | -34.6785 | -23.1741
13.8323 | -34.6791 | -23.1828
13.8384 | -34.6796 | -23.1912
13.8445 | -34.6802 | -23.1994
13.8506 | -34.6807 | -23.2076
13.8567 | -34.6813 | -23.2162
13.8628 | -34.6818 | -23.2251
13.8689 | -34.6824 | -23.2336
13.875 | -34.683 | -23.2425
13.8811 | -34.6835 | -23.2513
13.8872 | -34.6841 | -23.26
13.8933 | -34.6845 | -23.2686
13.8994 | -34.6849 | -23.2775
13.9055 | -34.6854 | -23.2862
13.9116 | -34.6858 | -23.2948
13.9177 | -34.6863 | -23.3036
13.9238 | -34.6867 | -23.3121
13.9299 | -34.6873 | -23.3203
13.936 | -34.6877 | -23.3284
13.9421 | -34.6882 | -23.337
13.9482 | -34.6886 | -23.3454
13.9543 | -34.6891 | -23.3541
13.9604 | -34.6896 | -23.3627
13.9665 | -34.6901 | -23.3711
13.9727 | -34.6906 | -23.3795
13.9788 | -34.6912 | -23.3884
13.9849 | -34.6917 | -23.397
13.991 | -34.6922 | -23.4058
13.9971 | -34.6927 | -23.4142
14.0032 | -34.6932 | -23.4225
14.0093 | -34.6937 | -23.4311
14.0154 | -34.6942 | -23.4395
14.0215 | -34.6947 | -23.4481
14.0276 | -34.6952 | -23.4568
14.0337 | -34.6958 | -23.4655
14.0398 | -34.6963 | -23.4738
14.0459 | -34.6968 | -23.4823
14.052 | -34.6972 | -23.4908
14.0581 | -34.6977 | -23.4996
14.0642 | -34.6982 | -23.508
14.0703 | -34.6986 | -23.5165
14.0764 | -34.6991 | -23.525
14.0825 | -34.6996 | -23.5333
14.0886 | -34.7001 | -23.5422
14.0947 | -34.7006 | -23.5508
14.1008 | -34.7011 | -23.5591
14.1069 | -34.7016 | -23.5673
14.1131 | -34.7021 | -23.5762
14.1192 | -34.7026 | -23.5846
14.1253 | -34.7031 | -23.5929
14.1314 | -34.7036 | -23.6013
14.1375 | -34.7041 | -23.6093
14.1436 | -34.7046 | -23.6178
14.1497 | -34.7051 | -23.6264
14.1558 | -34.7055 | -23.6348
14.1619 | -34.706 | -23.6433
14.168 | -34.7065 | -23.6521
14.1741 | -34.707 | -23.6606
14.1802 | -34.7075 | -23.669
14.1863 | -34.708 | -23.6776
14.1924 | -34.7085 | -23.6861
14.1985 | -34.709 | -23.6945
14.2046 | -34.7095 | -23.7029
14.2107 | -34.71 | -23.7113
14.2168 | -34.7104 | -23.7197
14.2229 | -34.7109 | -23.7281
14.229 | -34.7114 | -23.7369
14.2351 | -34.7119 | -23.7456
14.2412 | -34.7124 | -23.7542
14.2473 | -34.7129 | -23.7626
14.2534 | -34.7134 | -23.7709
14.2596 | -34.7139 | -23.7793
14.2657 | -34.7144 | -23.7876
14.2718 | -34.7149 | -23.796
14.2779 | -34.7154 | -23.8045
14.284 | -34.7159 | -23.8129
14.2901 | -34.7164 | -23.8213
14.2962 | -34.7169 | -23.8298
14.3023 | -34.7174 | -23.8384
14.3084 | -34.7179 | -23.8467
14.3145 | -34.7184 | -23.8553
14.3206 | -34.7188 | -23.8642
14.3267 | -34.7194 | -23.8728
14.3328 | -34.7199 | -23.8814
14.3389 | -34.7203 | -23.8898
14.345 | -34.7208 | -23.8983
14.3511 | -34.7213 | -23.9066
14.3572 | -34.7218 | -23.9152
14.3633 | -34.7223 | -23.9239
14.3694 | -34.7228 | -23.9323
14.3755 | -34.7233 | -23.9407
14.3816 | -34.7238 | -23.9492
14.3877 | -34.7243 | -23.9578
14.3938 | -34.7249 | -23.9661
14.4 | -34.7254 | -23.9746
14.4061 | -34.7258 | -23.9832
14.4122 | -34.7264 | -23.9916
14.4183 | -34.7269 | -24.0001
14.4244 | -34.7273 | -24.0087
14.4305 | -34.7278 | -24.0171
14.4366 | -34.7283 | -24.0256
14.4427 | -34.7288 | -24.0342
14.4488 | -34.7293 | -24.0425
14.4549 | -34.7298 | -24.0509
14.461 | -34.7302 | -24.0592
14.4671 | -34.7307 | -24.068
14.4732 | -34.7312 | -24.0766
14.4793 | -34.7317 | -24.085
14.4854 | -34.7322 | -24.0934
14.4915 | -34.7326 | -24.1017
14.4976 | -34.7331 | -24.1102
14.5037 | -34.7336 | -24.1186
14.5098 | -34.734 | -24.127
14.5159 | -34.7345 | -24.1354
14.522 | -34.7349 | -24.1438
14.5281 | -34.7354 | -24.1522
14.5342 | -34.7359 | -24.1606
14.5403 | -34.7364 | -24.1691
14.5465 | -34.7368 | -24.1777
14.5526 | -34.7373 | -24.1863
14.5587 | -34.7378 | -24.1947
14.5648 | -34.7383 | -24.2033
14.5709 | -34.7388 | -24.2116
14.577 | -34.7393 | -24.2201
14.5831 | -34.7397 | -24.2285
14.5892 | -34.7403 | -24.2368
14.5953 | -34.7407 | -24.245
14.6014 | -34.7412 | -24.2534
14.6075 | -34.7417 | -24.2618
14.6136 | -34.7421 | -24.2701
14.6197 | -34.7426 | -24.2786
14.6258 | -34.743 | -24.287
14.6319 | -34.7435 | -24.2952
14.638 | -34.7439 | -24.3037
14.6441 | -34.7444 | -24.3121
14.6502 | -34.7448 | -24.3207
14.6563 | -34.7453 | -24.3291
14.6624 | -34.7458 | -24.3376
14.6685 | -34.7463 | -24.346
14.6746 | -34.7467 | -24.3544
14.6807 | -34.7472 | -24.3627
14.6869 | -34.7477 | -24.3713
14.693 | -34.7482 | -24.3797
14.6991 | -34.7487 | -24.388
14.7052 | -34.7492 | -24.3964
14.7113 | -34.7497 | -24.4051
14.7174 | -34.7502 | -24.4138
14.7235 | -34.7507 | -24.422
14.7296 | -34.7512 | -24.4305
14.7357 | -34.7517 | -24.4391
14.7418 | -34.7522 | -24.4473
14.7479 | -34.7527 | -24.4555
14.754 | -34.7532 | -24.4637
14.7601 | -34.7536 | -24.4721
14.7662 | -34.7541 | -24.4805
14.7723 | -34.7546 | -24.4886
14.7784 | -34.755 | -24.4971
14.7845 | -34.7555 | -24.5056
14.7906 | -34.756 | -24.5143
14.7967 | -34.7564 | -24.5227
14.8028 | -34.7569 | -24.531
14.8089 | -34.7574 | -24.5394
14.815 | -34.7579 | -24.5478
14.8211 | -34.7584 | -24.5561
14.8272 | -34.7588 | -24.5645
14.8334 | -34.7593 | -24.573
14.8395 | -34.7598 | -24.5813
14.8456 | -34.7603 | -24.5893
14.8517 | -34.7608 | -24.5979
14.8578 | -34.7613 | -24.6062
14.8639 | -34.7617 | -24.6146
14.87 | -34.7622 | -24.6232
14.8761 | -34.7627 | -24.6316
14.8822 | -34.7632 | -24.64
14.8883 | -34.7636 | -24.6483
14.8944 | -34.7641 | -24.6566
14.9005 | -34.7646 | -24.6651
14.9066 | -34.7651 | -24.6734
14.9127 | -34.7656 | -24.6817
14.9188 | -34.766 | -24.6899
14.9249 | -34.7665 | -24.6981
14.931 | -34.767 | -24.7067
14.9371 | -34.7675 | -24.7153
14.9432 | -34.768 | -24.7239
14.9493 | -34.7684 | -24.7321
14.9554 | -34.7689 | -24.7405
14.9615 | -34.7694 | -24.7491
14.9676 | -34.7699 | -24.7573
14.9738 | -34.7704 | -24.7657
14.9799 | -34.7709 | -24.7743
14.986 | -34.7713 | -24.7827
14.9921 | -34.7718 | -24.791
14.9982 | -34.7722 | -24.7993
15.0043 | -34.7727 | -24.8077
15.0104 | -34.7732 | -24.8163
15.0165 | -34.7737 | -24.8248
15.0226 | -34.7742 | -24.8329
15.0287 | -34.7746 | -24.8412
15.0348 | -34.7751 | -24.8497
15.0409 | -34.7756 | -24.8583
15.047 | -34.7761 | -24.8667
15.0531 | -34.7766 | -24.8748
15.0592 | -34.7771 | -24.883
15.0653 | -34.7775 | -24.8914
15.0714 | -34.778 | -24.8997
15.0775 | -34.7785 | -24.9081
15.0836 | -34.779 | -24.9166
15.0897 | -34.7795 | -24.9247
15.0958 | -34.78 | -24.933
15.1019 | -34.7804 | -24.9413
15.108 | -34.7809 | -24.9495
15.1141 | -34.7814 | -24.9579
15.1203 | -34.7819 | -24.9662
15.1264 | -34.7824 | -24.9742
15.1325 | -34.7828 | -24.9828
15.1386 | -34.7833 | -24.991
15.1447 | -34.7838 | -24.9993
15.1508 | -34.7843 | -25.0077
15.1569 | -34.7848 | -25.0161
15.163 | -34.7852 | -25.0238
15.1691 | -34.7857 | -25.0321
15.1752 | -34.7862 | -25.0404
15.1813 | -34.7866 | -25.0487
15.1874 | -34.7871 | -25.057
15.1935 | -34.7876 | -25.0652
15.1996 | -34.7881 | -25.0734
15.2057 | -34.7886 | -25.0818
15.2118 | -34.7891 | -25.0902
15.2179 | -34.7895 | -25.0986
15.224 | -34.79 | -25.1068
15.2301 | -34.7905 | -25.1148
15.2362 | -34.7909 | -25.1231
15.2423 | -34.7914 | -25.1312
15.2484 | -34.7919 | -25.1395
15.2545 | -34.7924 | -25.148
15.2607 | -34.7929 | -25.1563
15.2668 | -34.7934 | -25.1641
15.2729 | -34.7938 | -25.1723
15.279 | -34.7943 | -25.1809
15.2851 | -34.7948 | -25.1891
15.2912 | -34.7953 | -25.1974
15.2973 | -34.7958 | -25.2056
15.3034 | -34.7963 | -25.2139
15.3095 | -34.7967 | -25.2224
15.3156 | -34.7972 | -25.2306
15.3217 | -34.7977 | -25.2389
15.3278 | -34.7982 | -25.2473
15.3339 | -34.7987 | -25.2557
15.34 | -34.7991 | -25.2638
15.3461 | -34.7996 | -25.272
15.3522 | -34.8001 | -25.2805
15.3583 | -34.8005 | -25.2889
15.3644 | -34.801 | -25.2971
15.3705 | -34.8015 | -25.3053
15.3766 | -34.802 | -25.3136
15.3827 | -34.8024 | -25.3218
15.3888 | -34.8029 | -25.3302
15.3949 | -34.8034 | -25.3384
15.401 | -34.8038 | -25.3468
15.4072 | -34.8043 | -25.3548
15.4133 | -34.8048 | -25.3631
15.4194 | -34.8052 | -25.3714
15.4255 | -34.8057 | -25.3797
15.4316 | -34.8062 | -25.388
15.4377 | -34.8067 | -25.3964
15.4438 | -34.8071 | -25.4046
15.4499 | -34.8076 | -25.4127
15.456 | -34.8081 | -25.4208
15.4621 | -34.8086 | -25.429
15.4682 | -34.809 | -25.4373
15.4743 | -34.8095 | -25.4453
15.4804 | -34.81 | -25.4534
15.4865 | -34.8105 | -25.4614
15.4926 | -34.8109 | -25.4696
15.4987 | -34.8114 | -25.478
15.5048 | -34.8119 | -25.486
15.5109 | -34.8123 | -25.4941
15.517 | -34.8128 | -25.5023
15.5231 | -34.8133 | -25.5105
15.5292 | -34.8137 | -25.5186
15.5353 | -34.8142 | -25.527
15.5414 | -34.8147 | -25.5351
15.5476 | -34.8152 | -25.543
15.5537 | -34.8157 | -25.5512
15.5598 | -34.8162 | -25.5595
15.5659 | -34.8167 | -25.5674
15.572 | -34.8172 | -25.5756
15.5781 | -34.8176 | -25.5836
15.5842 | -34.8181 | -25.5915
15.5903 | -34.8186 | -25.5994
15.5964 | -34.8191 | -25.6073
15.6025 | -34.8196 | -25.6153
15.6086 | -34.82 | -25.6236
15.6147 | -34.8205 | -25.6318
15.6208 | -34.821 | -25.6396
15.6269 | -34.8215 | -25.6476
15.633 | -34.822 | -25.6558
15.6391 | -34.8225 | -25.664
15.6452 | -34.8229 | -25.6724
15.6513 | -34.8234 | -25.6803
15.6574 | -34.8239 | -25.6885
15.6635 | -34.8244 | -25.6966
15.6696 | -34.8249 | -25.7048
15.6757 | -34.8254 | -25.7131
15.6818 | -34.8259 | -25.7214
15.688 | -34.8264 | -25.7293
15.6941 | -34.8269 | -25.7374
15.7002 | -34.8274 | -25.7456
15.7063 | -34.8278 | -25.7537
15.7124 | -34.8283 | -25.762
15.7185 | -34.8288 | -25.7701
15.7246 | -34.8293 | -25.7783
15.7307 | -34.8298 | -25.7863
15.7368 | -34.8303 | -25.7945
15.7429 | -34.8308 | -25.8027
15.749 | -34.8312 | -25.8109
15.7551 | -34.8317 | -25.8191
15.7612 | -34.8322 | -25.8271
15.7673 | -34.8327 | -25.835
15.7734 | -34.8332 | -25.8432
15.7795 | -34.8337 | -25.8513
15.7856 | -34.8342 | -25.8594
15.7917 | -34.8347 | -25.8674
15.7978 | -34.8351 | -25.8757
15.8039 | -34.8356 | -25.8837
15.81 | -34.8361 | -25.8919
15.8161 | -34.8365 | -25.8999
15.8222 | -34.837 | -25.9078
15.8283 | -34.8375 | -25.9157
15.8345 | -34.838 | -25.9239
15.8406 | -34.8385 | -25.9319
15.8467 | -34.839 | -25.94
15.8528 | -34.8395 | -25.9482
15.8589 | -34.84 | -25.9563
15.865 | -34.8406 | -25.9639
15.8711 | -34.841 | -25.9721
15.8772 | -34.8415 | -25.9802
15.8833 | -34.842 | -25.9885
15.8894 | -34.8425 | -25.9966
15.8955 | -34.8429 | -26.0047
15.9016 | -34.8434 | -26.0126
15.9077 | -34.8439 | -26.0208
15.9138 | -34.8444 | -26.0289
15.9199 | -34.8449 | -26.0369
15.926 | -34.8453 | -26.0449
15.9321 | -34.8459 | -26.0529
15.9382 | -34.8463 | -26.0609
15.9443 | -34.8468 | -26.069
15.9504 | -34.8473 | -26.0771
15.9565 | -34.8478 | -26.0852
15.9626 | -34.8483 | -26.0933
15.9687 | -34.8488 | -26.1013
15.9749 | -34.8493 | -26.1092
15.981 | -34.8497 | -26.1173
15.9871 | -34.8502 | -26.1254
15.9932 | -34.8507 | -26.1333
15.9993 | -34.8512 | -26.1414
16.0054 | -34.8517 | -26.149
16.0115 | -34.8522 | -26.1571
16.0176 | -34.8527 | -26.1653
16.0237 | -34.8532 | -26.1734
16.0298 | -34.8536 | -26.1815
16.0359 | -34.8541 | -26.1894
16.042 | -34.8547 | -26.1971
16.0481 | -34.8551 | -26.2052
16.0542 | -34.8556 | -26.2134
16.0603 | -34.8561 | -26.2213
16.0664 | -34.8566 | -26.2293
16.0725 | -34.8571 | -26.2371
16.0786 | -34.8576 | -26.245
16.0847 | -34.8581 | -26.2531
16.0908 | -34.8586 | -26.2611
16.0969 | -34.8591 | -26.2693
16.103 | -34.8596 | -26.2771
16.1091 | -34.8601 | -26.2848
16.1152 | -34.8606 | -26.2928
16.1214 | -34.8611 | -26.301
16.1275 | -34.8616 | -26.3087
16.1336 | -34.8621 | -26.3168
16.1397 | -34.8626 | -26.3247
16.1458 | -34.8631 | -26.3326
16.1519 | -34.8636 | -26.3407
16.158 | -34.864 | -26.3488
16.1641 | -34.8645 | -26.3568
16.1702 | -34.865 | -26.3649
16.1763 | -34.8655 | -26.3729
16.1824 | -34.866 | -26.3806
16.1885 | -34.8665 | -26.3885
16.1946 | -34.867 | -26.3964
16.2007 | -34.8675 | -26.4044
16.2068 | -34.868 | -26.4122
16.2129 | -34.8685 | -26.4203
16.219 | -34.869 | -26.428
16.2251 | -34.8695 | -26.4359
16.2312 | -34.87 | -26.4438
16.2373 | -34.8705 | -26.4517
16.2434 | -34.871 | -26.4596
16.2495 | -34.8715 | -26.4675
16.2556 | -34.872 | -26.4755
16.2618 | -34.8725 | -26.4833
16.2679 | -34.873 | -26.4913
16.274 | -34.8735 | -26.4992
16.2801 | -34.874 | -26.5071
16.2862 | -34.8745 | -26.5151
16.2923 | -34.875 | -26.523
16.2984 | -34.8755 | -26.5311
16.3045 | -34.876 | -26.5391
16.3106 | -34.8765 | -26.5471
16.3167 | -34.8769 | -26.555
16.3228 | -34.8775 | -26.5628
16.3289 | -34.8779 | -26.5708
16.335 | -34.8785 | -26.5787
16.3411 | -34.8789 | -26.5867
16.3472 | -34.8794 | -26.5946
16.3533 | -34.88 | -26.6026
16.3594 | -34.8805 | -26.6102
16.3655 | -34.8809 | -26.6182
16.3716 | -34.8814 | -26.626
16.3777 | -34.8819 | -26.6339
16.3838 | -34.8824 | -26.6418
16.3899 | -34.8829 | -26.6497
16.396 | -34.8834 | -26.6576
16.4021 | -34.8839 | -26.6655
16.4083 | -34.8844 | -26.6734
16.4144 | -34.8849 | -26.6813
16.4205 | -34.8854 | -26.6892
16.4266 | -34.8859 | -26.6968
16.4327 | -34.8864 | -26.7046
16.4388 | -34.8868 | -26.7125
16.4449 | -34.8874 | -26.7204
16.451 | -34.8878 | -26.7283
16.4571 | -34.8883 | -26.7363
16.4632 | -34.8888 | -26.744
16.4693 | -34.8893 | -26.7519
16.4754 | -34.8898 | -26.7599
16.4815 | -34.8903 | -26.7677
16.4876 | -34.8908 | -26.7755
16.4937 | -34.8913 | -26.7834
16.4998 | -34.8919 | -26.7912
16.5059 | -34.8924 | -26.799
16.512 | -34.8929 | -26.807
16.5181 | -34.8934 | -26.815
16.5242 | -34.8939 | -26.823
16.5303 | -34.8944 | -26.8307
16.5364 | -34.8948 | -26.8385
16.5425 | -34.8953 | -26.8466
16.5487 | -34.8958 | -26.8545
16.5548 | -34.8963 | -26.8623
16.5609 | -34.8968 | -26.8702
16.567 | -34.8974 | -26.878
16.5731 | -34.8979 | -26.8859
16.5792 | -34.8983 | -26.8937
16.5853 | -34.8989 | -26.9015
16.5914 | -34.8994 | -26.9095
16.5975 | -34.8999 | -26.9174
16.6036 | -34.9004 | -26.9251
16.6097 | -34.9009 | -26.933
16.6158 | -34.9013 | -26.9408
16.6219 | -34.9019 | -26.9487
16.628 | -34.9024 | -26.9566
16.6341 | -34.9029 | -26.9645
16.6402 | -34.9034 | -26.9722
16.6463 | -34.9039 | -26.9801
16.6524 | -34.9044 | -26.988
16.6585 | -34.9049 | -26.9961
16.6646 | -34.9054 | -27.0039
16.6707 | -34.9059 | -27.0115
16.6768 | -34.9064 | -27.0193
16.6829 | -34.9069 | -27.0271
16.689 | -34.9074 | -27.035
16.6952 | -34.9079 | -27.0428
16.7013 | -34.9084 | -27.0505
16.7074 | -34.9089 | -27.0581
16.7135 | -34.9094 | -27.0658
16.7196 | -34.9099 | -27.0737
16.7257 | -34.9104 | -27.0816
16.7318 | -34.9109 | -27.0894
16.7379 | -34.9114 | -27.0973
16.744 | -34.9119 | -27.1048
16.7501 | -34.9124 | -27.1127
16.7562 | -34.9129 | -27.1206
16.7623 | -34.9134 | -27.1283
16.7684 | -34.9139 | -27.1362
16.7745 | -34.9144 | -27.1441
16.7806 | -34.915 | -27.1516
16.7867 | -34.9155 | -27.1594
16.7928 | -34.916 | -27.1671
16.7989 | -34.9165 | -27.1748
16.805 | -34.917 | -27.1826
16.8111 | -34.9175 | -27.1904
16.8172 | -34.918 | -27.1982
16.8233 | -34.9186 | -27.206
16.8294 | -34.9191 | -27.2138
16.8356 | -34.9196 | -27.2217
16.8417 | -34.92 | -27.2296
16.8478 | -34.9206 | -27.2373
16.8539 | -34.9211 | -27.245
16.86 | -34.9216 | -27.2528
16.8661 | -34.9221 | -27.2608
16.8722 | -34.9226 | -27.2688
16.8783 | -34.9231 | -27.2767
16.8844 | -34.9237 | -27.2843
16.8905 | -34.9242 | -27.292
16.8966 | -34.9246 | -27.2999
16.9027 | -34.9252 | -27.3077
16.9088 | -34.9256 | -27.3156
16.9149 | -34.9262 | -27.3234
16.921 | -34.9267 | -27.331
16.9271 | -34.9272 | -27.3391
16.9332 | -34.9277 | -27.3469
16.9393 | -34.9282 | -27.3547
16.9454 | -34.9287 | -27.3623
16.9515 | -34.9292 | -27.3702
16.9576 | -34.9297 | -27.3778
16.9637 | -34.9302 | -27.3854
16.9698 | -34.9307 | -27.393
16.9759 | -34.9312 | -27.4008
16.9821 | -34.9317 | -27.4086
16.9882 | -34.9323 | -27.4161
16.9943 | -34.9328 | -27.4238
17.0004 | -34.9333 | -27.4315
17.0065 | -34.9338 | -27.4392
17.0126 | -34.9343 | -27.447
17.0187 | -34.9348 | -27.4549
17.0248 | -34.9353 | -27.4624
17.0309 | -34.9358 | -27.4701
17.037 | -34.9363 | -27.4778
17.0431 | -34.9368 | -27.4855
17.0492 | -34.9373 | -27.4933
17.0553 | -34.9378 | -27.501
17.0614 | -34.9383 | -27.5087
17.0675 | -34.9388 | -27.5165
17.0736 | -34.9394 | -27.5243
17.0797 | -34.9399 | -27.532
17.0858 | -34.9404 | -27.5398
17.0919 | -34.9409 | -27.5476
17.098 | -34.9414 | -27.5551
17.1041 | -34.9419 | -27.5628
17.1102 | -34.9425 | -27.5705
17.1163 | -34.943 | -27.5782
17.1225 | -34.9435 | -27.5858
17.1286 | -34.944 | -27.5934
17.1347 | -34.9445 | -27.6012
17.1408 | -34.9451 | -27.6091
17.1469 | -34.9456 | -27.6168
17.153 | -34.9461 | -27.6245
17.1591 | -34.9466 | -27.6321
17.1652 | -34.9471 | -27.6397
17.1713 | -34.9476 | -27.6474
17.1774 | -34.9481 | -27.655
17.1835 | -34.9486 | -27.6626
17.1896 | -34.9491 | -27.6702
17.1957 | -34.9496 | -27.678
17.2018 | -34.9501 | -27.6854
17.2079 | -34.9506 | -27.6931
17.214 | -34.9511 | -27.7009
17.2201 | -34.9516 | -27.7084
17.2262 | -34.9521 | -27.7162
17.2323 | -34.9527 | -27.7239
17.2384 | -34.9532 | -27.7314
17.2445 | -34.9537 | -27.7391
17.2506 | -34.9542 | -27.7468
17.2567 | -34.9547 | -27.7545
17.2628 | -34.9552 | -27.7622
17.269 | -34.9557 | -27.7699
17.2751 | -34.9562 | -27.7776
17.2812 | -34.9567 | -27.7854
17.2873 | -34.9572 | -27.793
17.2934 | -34.9577 | -27.8006
17.2995 | -34.9582 | -27.8082
17.3056 | -34.9588 | -27.8158
17.3117 | -34.9592 | -27.8235
17.3178 | -34.9598 | -27.8311
17.3239 | -34.9603 | -27.8388
17.33 | -34.9607 | -27.8465
17.3361 | -34.9612 | -27.8542
17.3422 | -34.9617 | -27.8616
17.3483 | -34.9622 | -27.8695
17.3544 | -34.9627 | -27.8771
17.3605 | -34.9632 | -27.8848
17.3666 | -34.9637 | -27.8923
17.3727 | -34.9642 | -27.9
17.3788 | -34.9647 | -27.9075
17.3849 | -34.9652 | -27.9154
17.391 | -34.9657 | -27.923
17.3971 | -34.9662 | -27.9307
17.4032 | -34.9667 | -27.9384
17.4094 | -34.9672 | -27.946
17.4155 | -34.9677 | -27.9535
17.4216 | -34.9682 | -27.9611
17.4277 | -34.9687 | -27.9689
17.4338 | -34.9692 | -27.9765
17.4399 | -34.9697 | -27.9842
17.446 | -34.9702 | -27.9918
17.4521 | -34.9708 | -27.9992
17.4582 | -34.9713 | -28.0069
17.4643 | -34.9718 | -28.0145
17.4704 | -34.9723 | -28.022
17.4765 | -34.9728 | -28.0298
17.4826 | -34.9733 | -28.0374
17.4887 | -34.9738 | -28.0452
17.4948 | -34.9743 | -28.0528
17.5009 | -34.9748 | -28.0605
17.507 | -34.9753 | -28.0683
17.5131 | -34.9758 | -28.0758
17.5192 | -34.9763 | -28.0834
17.5253 | -34.9769 | -28.0911
17.5314 | -34.9774 | -28.0987
17.5375 | -34.9778 | -28.1063
17.5436 | -34.9784 | -28.1141
17.5497 | -34.9789 | -28.1217
17.5559 | -34.9794 | -28.1292
17.562 | -34.9799 | -28.1369
17.5681 | -34.9804 | -28.1445
17.5742 | -34.9809 | -28.1522
17.5803 | -34.9815 | -28.1598
17.5864 | -34.982 | -28.1672
17.5925 | -34.9825 | -28.1748
17.5986 | -34.983 | -28.1825
17.6047 | -34.9835 | -28.1903
17.6108 | -34.984 | -28.198
17.6169 | -34.9845 | -28.2054
17.623 | -34.985 | -28.2129
17.6291 | -34.9855 | -28.2205
17.6352 | -34.986 | -28.228
17.6413 | -34.9865 | -28.2357
17.6474 | -34.9871 | -28.2432
17.6535 | -34.9876 | -28.2507
17.6596 | -34.9881 | -28.2583
17.6657 | -34.9886 | -28.266
17.6718 | -34.9891 | -28.2735
17.6779 | -34.9896 | -28.2809
17.684 | -34.9901 | -28.2885
17.6901 | -34.9906 | -28.296
17.6963 | -34.9911 | -28.3036
17.7024 | -34.9916 | -28.311
17.7085 | -34.9921 | -28.3185
17.7146 | -34.9926 | -28.3261
17.7207 | -34.9931 | -28.3337
17.7268 | -34.9936 | -28.3413
17.7329 | -34.9941 | -28.3491
17.739 | -34.9946 | -28.3568
17.7451 | -34.9951 | -28.3643
17.7512 | -34.9956 | -28.3719
17.7573 | -34.9961 | -28.3796
17.7634 | -34.9966 | -28.387
17.7695 | -34.9971 | -28.3947
17.7756 | -34.9976 | -28.4021
17.7817 | -34.9981 | -28.4097
17.7878 | -34.9987 | -28.4173
17.7939 | -34.9992 | -28.4249
17.8 | -34.9997 | -28.4323
17.8061 | -35.0002 | -28.4399
17.8122 | -35.0007 | -28.4475
17.8183 | -35.0013 | -28.4549
17.8244 | -35.0018 | -28.4626
17.8305 | -35.0023 | -28.4701
17.8366 | -35.0028 | -28.4775
17.8428 | -35.0033 | -28.4851
17.8489 | -35.0038 | -28.4927
17.855 | -35.0043 | -28.5002
17.8611 | -35.0048 | -28.5078
17.8672 | -35.0053 | -28.5152
17.8733 | -35.0058 | -28.5226
17.8794 | -35.0063 | -28.5299
17.8855 | -35.0068 | -28.5374
17.8916 | -35.0073 | -28.5453
17.8977 | -35.0078 | -28.553
17.9038 | -35.0084 | -28.5603
17.9099 | -35.0089 | -28.568
17.916 | -35.0094 | -28.5756
17.9221 | -35.0099 | -28.5833
17.9282 | -35.0104 | -28.5908
17.9343 | -35.0109 | -28.5984
17.9404 | -35.0114 | -28.6058
17.9465 | -35.0119 | -28.6134
17.9526 | -35.0124 | -28.6211
17.9587 | -35.0129 | -28.6286
17.9648 | -35.0134 | -28.6363
17.9709 | -35.0139 | -28.644
17.977 | -35.0144 | -28.6516
17.9832 | -35.0149 | -28.6592
17.9893 | -35.0154 | -28.6665
17.9954 | -35.0159 | -28.6742
18.0015 | -35.0164 | -28.6817
18.0076 | -35.017 | -28.6891
18.0137 | -35.0175 | -28.6967
18.0198 | -35.018 | -28.7041
18.0259 | -35.0185 | -28.7118
18.032 | -35.019 | -28.7194
18.0381 | -35.0195 | -28.7271
18.0442 | -35.02 | -28.7345
18.0503 | -35.0205 | -28.742
18.0564 | -35.021 | -28.7495
18.0625 | -35.0215 | -28.7571
18.0686 | -35.022 | -28.7647
18.0747 | -35.0225 | -28.7723
18.0808 | -35.023 | -28.7797
18.0869 | -35.0235 | -28.7873
18.093 | -35.0241 | -28.7949
18.0991 | -35.0245 | -28.8025
18.1052 | -35.0251 | -28.8101
18.1113 | -35.0255 | -28.8175
18.1174 | -35.026 | -28.8248
18.1236 | -35.0265 | -28.8325
18.1297 | -35.027 | -28.8401
18.1358 | -35.0275 | -28.8474
18.1419 | -35.028 | -28.855
18.148 | -35.0285 | -28.8623
18.1541 | -35.029 | -28.8696
18.1602 | -35.0295 | -28.8772
18.1663 | -35.03 | -28.8847
18.1724 | -35.0305 | -28.8922
18.1785 | -35.031 | -28.8996
18.1846 | -35.0315 | -28.907
18.1907 | -35.0321 | -28.9143
18.1968 | -35.0325 | -28.9219
18.2029 | -35.0331 | -28.9296
18.209 | -35.0335 | -28.9372
18.2151 | -35.034 | -28.9452
18.2212 | -35.0345 | -28.9526
18.2273 | -35.0351 | -28.9603
18.2334 | -35.0356 | -28.9679
18.2395 | -35.036 | -28.9751
18.2456 | -35.0366 | -28.9824
18.2517 | -35.0371 | -28.99
18.2578 | -35.0375 | -28.9974
18.2639 | -35.0381 | -29.005
18.2701 | -35.0386 | -29.0126
18.2762 | -35.0391 | -29.0199
18.2823 | -35.0396 | -29.0272
18.2884 | -35.0401 | -29.0346
18.2945 | -35.0406 | -29.0418
18.3006 | -35.0411 | -29.0492
18.3067 | -35.0416 | -29.0569
18.3128 | -35.0421 | -29.0644
18.3189 | -35.0426 | -29.072
18.325 | -35.0431 | -29.0794
18.3311 | -35.0436 | -29.0869
18.3372 | -35.0441 | -29.0946
18.3433 | -35.0446 | -29.102
18.3494 | -35.0451 | -29.1095
18.3555 | -35.0456 | -29.117
18.3616 | -35.0462 | -29.1242
18.3677 | -35.0467 | -29.1318
18.3738 | -35.0472 | -29.1391
18.3799 | -35.0477 | -29.1469
18.386 | -35.0482 | -29.1542
18.3921 | -35.0486 | -29.1618
18.3982 | -35.0491 | -29.1691
18.4043 | -35.0496 | -29.1765
18.4105 | -35.0501 | -29.184
18.4166 | -35.0506 | -29.1914
18.4227 | -35.0511 | -29.1989
18.4288 | -35.0516 | -29.2064
18.4349 | -35.0521 | -29.2138
18.441 | -35.0525 | -29.2212
18.4471 | -35.053 | -29.2285
18.4532 | -35.0535 | -29.2361
18.4593 | -35.054 | -29.2436
18.4654 | -35.0545 | -29.2509
18.4715 | -35.055 | -29.2587
18.4776 | -35.0555 | -29.2663
18.4837 | -35.056 | -29.2737
18.4898 | -35.0565 | -29.2811
18.4959 | -35.057 | -29.2887
18.502 | -35.0575 | -29.2961
18.5081 | -35.058 | -29.3036
18.5142 | -35.0585 | -29.3109
18.5203 | -35.059 | -29.3184
18.5264 | -35.0595 | -29.3259
18.5325 | -35.06 | -29.3333
18.5386 | -35.0605 | -29.3408
18.5447 | -35.0609 | -29.3482
18.5508 | -35.0615 | -29.3557
18.557 | -35.0619 | -29.3632
18.5631 | -35.0624 | -29.3705
18.5692 | -35.0629 | -29.3776
18.5753 | -35.0634 | -29.3851
18.5814 | -35.0639 | -29.3924
18.5875 | -35.0644 | -29.3996
18.5936 | -35.0649 | -29.407
18.5997 | -35.0653 | -29.4148
18.6058 | -35.0658 | -29.422
18.6119 | -35.0663 | -29.4291
18.618 | -35.0668 | -29.4368
18.6241 | -35.0673 | -29.4443
18.6302 | -35.0678 | -29.4522
18.6363 | -35.0683 | -29.4597
18.6424 | -35.0688 | -29.4669
18.6485 | -35.0692 | -29.4746
18.6546 | -35.0697 | -29.482
18.6607 | -35.0702 | -29.4896
18.6668 | -35.0707 | -29.4967
18.6729 | -35.0712 | -29.5041
18.679 | -35.0717 | -29.5114
18.6851 | -35.0722 | -29.5188
18.6912 | -35.0727 | -29.5262
18.6974 | -35.0731 | -29.5336
18.7035 | -35.0736 | -29.5409
18.7096 | -35.0741 | -29.5483
18.7157 | -35.0746 | -29.5556
18.7218 | -35.0751 | -29.5629
18.7279 | -35.0756 | -29.5701
18.734 | -35.0761 | -29.5775
18.7401 | -35.0766 | -29.585
18.7462 | -35.0771 | -29.5921
18.7523 | -35.0776 | -29.5997
18.7584 | -35.0782 | -29.6072
18.7645 | -35.0787 | -29.6146
18.7706 | -35.0792 | -29.6221
18.7767 | -35.0797 | -29.6298
18.7828 | -35.0801 | -29.6367
18.7889 | -35.0806 | -29.644
18.795 | -35.0811 | -29.6518
18.8011 | -35.0816 | -29.6592
18.8072 | -35.0821 | -29.6668
18.8133 | -35.0826 | -29.6742
18.8194 | -35.0831 | -29.6814
18.8255 | -35.0836 | -29.6888
18.8316 | -35.0841 | -29.6966
18.8377 | -35.0846 | -29.704
18.8439 | -35.0851 | -29.7114
18.85 | -35.0856 | -29.7188
18.8561 | -35.0861 | -29.7262
18.8622 | -35.0866 | -29.7335
18.8683 | -35.0871 | -29.7407
18.8744 | -35.0876 | -29.7478
18.8805 | -35.0881 | -29.7551
18.8866 | -35.0886 | -29.7626
18.8927 | -35.0891 | -29.7698
18.8988 | -35.0895 | -29.7768
18.9049 | -35.09 | -29.7842
18.911 | -35.0905 | -29.7918
18.9171 | -35.091 | -29.7994
18.9232 | -35.0914 | -29.8064
18.9293 | -35.0919 | -29.814
18.9354 | -35.0924 | -29.8211
18.9415 | -35.0929 | -29.8288
18.9476 | -35.0934 | -29.836
18.9537 | -35.0938 | -29.8433
18.9598 | -35.0944 | -29.8505
18.9659 | -35.0949 | -29.858
18.972 | -35.0954 | -29.8654
18.9781 | -35.0959 | -29.8731
18.9843 | -35.0964 | -29.8807
18.9904 | -35.0969 | -29.8883
18.9965 | -35.0974 | -29.8954
19.0026 | -35.0979 | -29.9025
19.0087 | -35.0984 | -29.91
19.0148 | -35.0989 | -29.9171
19.0209 | -35.0994 | -29.9245
19.027 | -35.0999 | -29.9321
19.0331 | -35.1004 | -29.9393
19.0392 | -35.1009 | -29.9465
19.0453 | -35.1014 | -29.9537
19.0514 | -35.1019 | -29.9613
19.0575 | -35.1024 | -29.9689
19.0636 | -35.1029 | -29.9762
19.0697 | -35.1034 | -29.9834
19.0758 | -35.1038 | -29.991
19.0819 | -35.1043 | -29.9981
19.088 | -35.1048 | -30.0053
19.0941 | -35.1052 | -30.0124
19.1002 | -35.1057 | -30.0197
19.1063 | -35.1062 | -30.0269
19.1124 | -35.1067 | -30.034
19.1185 | -35.1072 | -30.0415
19.1246 | -35.1077 | -30.0489
19.1308 | -35.1082 | -30.0563
19.1369 | -35.1087 | -30.0635
19.143 | -35.1092 | -30.0711
19.1491 | -35.1096 | -30.0783
19.1552 | -35.1101 | -30.0855
19.1613 | -35.1106 | -30.093
19.1674 | -35.1111 | -30.1001
19.1735 | -35.1116 | -30.1072
19.1796 | -35.1121 | -30.1148
19.1857 | -35.1125 | -30.1222
19.1918 | -35.1129 | -30.1295
19.1979 | -35.1134 | -30.137
19.204 | -35.1139 | -30.1441
19.2101 | -35.1145 | -30.1511
19.2162 | -35.115 | -30.1583
19.2223 | -35.1155 | -30.1657
19.2284 | -35.1159 | -30.173
19.2345 | -35.1164 | -30.1805
19.2406 | -35.1169 | -30.1877
19.2467 | -35.1173 | -30.1949
19.2528 | -35.1178 | -30.2022
19.2589 | -35.1182 | -30.2096
19.265 | -35.1187 | -30.2168
19.2712 | -35.1192 | -30.2242
19.2773 | -35.1197 | -30.2315
19.2834 | -35.1202 | -30.2387
19.2895 | -35.1207 | -30.2461
19.2956 | -35.1211 | -30.2533
19.3017 | -35.1216 | -30.2605
19.3078 | -35.1221 | -30.2679
19.3139 | -35.1226 | -30.2747
19.32 | -35.1231 | -30.2819
19.3261 | -35.1235 | -30.2892
19.3322 | -35.124 | -30.2966
19.3383 | -35.1245 | -30.3039
19.3444 | -35.125 | -30.3107
19.3505 | -35.1254 | -30.3181
19.3566 | -35.126 | -30.3252
19.3627 | -35.1264 | -30.3326
19.3688 | -35.1269 | -30.3398
19.3749 | -35.1274 | -30.3469
19.381 | -35.1279 | -30.3538
19.3871 | -35.1284 | -30.361
19.3932 | -35.1288 | -30.3681
19.3993 | -35.1293 | -30.3755
19.4054 | -35.1298 | -30.3824
19.4115 | -35.1302 | -30.3896
19.4177 | -35.1307 | -30.3965
19.4238 | -35.1313 | -30.4037
19.4299 | -35.1318 | -30.4112
19.436 | -35.1323 | -30.4185
19.4421 | -35.1328 | -30.4256
19.4482 | -35.1333 | -30.433
19.4543 | -35.1338 | -30.4398
19.4604 | -35.1343 | -30.4469
19.4665 | -35.1347 | -30.454
19.4726 | -35.1352 | -30.4611
19.4787 | -35.1357 | -30.4684
19.4848 | -35.1362 | -30.4754
19.4909 | -35.1367 | -30.4826
19.497 | -35.1371 | -30.49
19.5031 | -35.1376 | -30.4973
19.5092 | -35.138 | -30.5044
19.5153 | -35.1385 | -30.5115
19.5214 | -35.139 | -30.5184
19.5275 | -35.1394 | -30.5255
19.5336 | -35.14 | -30.533
19.5397 | -35.1404 | -30.5403
19.5458 | -35.1409 | -30.5472
19.5519 | -35.1414 | -30.5544
19.5581 | -35.1418 | -30.5615
19.5642 | -35.1423 | -30.5684
19.5703 | -35.1428 | -30.5751
19.5764 | -35.1432 | -30.5823
19.5825 | -35.1438 | -30.5895
19.5886 | -35.1443 | -30.5966
19.5947 | -35.1448 | -30.6035
19.6008 | -35.1452 | -30.6101
19.6069 | -35.1458 | -30.6173
19.613 | -35.1463 | -30.6247
19.6191 | -35.1467 | -30.6314
19.6252 | -35.1473 | -30.6389
19.6313 | -35.1477 | -30.6456
19.6374 | -35.1481 | -30.653
19.6435 | -35.1486 | -30.6596
19.6496 | -35.1491 | -30.6664
19.6557 | -35.1495 | -30.6735
19.6618 | -35.15 | -30.6806
19.6679 | -35.1504 | -30.6881
19.674 | -35.1509 | -30.6954
19.6801 | -35.1513 | -30.7028
19.6862 | -35.1518 | -30.7097
19.6923 | -35.1522 | -30.7167
19.6984 | -35.1527 | -30.7236
19.7046 | -35.1531 | -30.7305
19.7107 | -35.1536 | -30.7376
19.7168 | -35.154 | -30.7447
19.7229 | -35.1545 | -30.7517
19.729 | -35.1549 | -30.7588
19.7351 | -35.1554 | -30.7656
19.7412 | -35.1559 | -30.7728
19.7473 | -35.1564 | -30.78
19.7534 | -35.157 | -30.7871
19.7595 | -35.1575 | -30.7941
19.7656 | -35.158 | -30.8013
19.7717 | -35.1585 | -30.808
19.7778 | -35.159 | -30.815
19.7839 | -35.1595 | -30.8222
19.79 | -35.16 | -30.8297
19.7961 | -35.1604 | -30.8367
19.8022 | -35.1608 | -30.8437
19.8083 | -35.1613 | -30.851
19.8144 | -35.1618 | -30.858
19.8205 | -35.1624 | -30.8652
19.8266 | -35.1628 | -30.8723
19.8327 | -35.1633 | -30.8791
19.8388 | -35.1638 | -30.8862
19.845 | -35.1642 | -30.8932
19.8511 | -35.1647 | -30.9005
19.8572 | -35.1651 | -30.9078
19.8633 | -35.1656 | -30.9145
19.8694 | -35.1661 | -30.9216
19.8755 | -35.1665 | -30.9286
19.8816 | -35.167 | -30.9356
19.8877 | -35.1674 | -30.9427
19.8938 | -35.1679 | -30.95
19.8999 | -35.1683 | -30.9574
19.906 | -35.1688 | -30.9646
19.9121 | -35.1693 | -30.9715
19.9182 | -35.1697 | -30.9786
19.9243 | -35.1702 | -30.9855
19.9304 | -35.1707 | -30.9925
19.9365 | -35.1711 | -30.9994
19.9426 | -35.1716 | -31.0062
19.9487 | -35.1721 | -31.0134
19.9548 | -35.1725 | -31.0207
19.9609 | -35.173 | -31.0278
19.967 | -35.1735 | -31.0348
19.9731 | -35.174 | -31.0418
19.9792 | -35.1745 | -31.0486
19.9853 | -35.175 | -31.0557
19.9915 | -35.1755 | -31.0623
19.9976 | -35.1759 | -31.0697
20.0037 | -35.1764 | -31.0769
20.0098 | -35.1768 | -31.0839
20.0159 | -35.1773 | -31.0907
20.022 | -35.1778 | -31.0975
20.0281 | -35.1782 | -31.1046
20.0342 | -35.1786 | -31.1115
20.0403 | -35.1792 | -31.1185
20.0464 | -35.1796 | -31.1257
20.0525 | -35.1801 | -31.1326
20.0586 | -35.1806 | -31.1396
20.0647 | -35.181 | -31.1466
20.0708 | -35.1815 | -31.1537
20.0769 | -35.1819 | -31.161
20.083 | -35.1824 | -31.1677
20.0891 | -35.1828 | -31.175
20.0952 | -35.1832 | -31.1823
20.1013 | -35.1837 | -31.1892
20.1074 | -35.1842 | -31.1961
20.1135 | -35.1847 | -31.2034
20.1196 | -35.1851 | -31.2103
20.1257 | -35.1856 | -31.2175
20.1319 | -35.1861 | -31.2243
20.138 | -35.1865 | -31.2312
20.1441 | -35.187 | -31.2385
20.1502 | -35.1874 | -31.2456
20.1563 | -35.1879 | -31.2524
20.1624 | -35.1884 | -31.2594
20.1685 | -35.1888 | -31.2666
20.1746 | -35.1893 | -31.2738
20.1807 | -35.1898 | -31.281
20.1868 | -35.1903 | -31.2881
20.1929 | -35.1908 | -31.295
20.199 | -35.1913 | -31.3018
20.2051 | -35.1918 | -31.3089
20.2112 | -35.1923 | -31.3161
20.2173 | -35.1927 | -31.3229
20.2234 | -35.1932 | -31.3297
20.2295 | -35.1936 | -31.3368
20.2356 | -35.1941 | -31.3441
20.2417 | -35.1946 | -31.3511
20.2478 | -35.1951 | -31.3581
20.2539 | -35.1955 | -31.3653
20.26 | -35.196 | -31.3718
20.2661 | -35.1964 | -31.3788
20.2723 | -35.1969 | -31.3858
20.2784 | -35.1974 | -31.3928
20.2845 | -35.1978 | -31.3997
20.2906 | -35.1983 | -31.4063
20.2967 | -35.1988 | -31.4132
20.3028 | -35.1993 | -31.4203
20.3089 | -35.1998 | -31.4273
20.315 | -35.2002 | -31.4341
20.3211 | -35.2007 | -31.4411
20.3272 | -35.2011 | -31.4481
20.3333 | -35.2016 | -31.4546
20.3394 | -35.2021 | -31.4616
20.3455 | -35.2025 | -31.4683
20.3516 | -35.203 | -31.4749
20.3577 | -35.2035 | -31.4816
20.3638 | -35.2039 | -31.4884
20.3699 | -35.2044 | -31.495
20.376 | -35.2049 | -31.502
20.3821 | -35.2054 | -31.5091
20.3882 | -35.2059 | -31.516
20.3943 | -35.2063 | -31.5232
20.4004 | -35.2068 | -31.53
20.4065 | -35.2072 | -31.5368
20.4126 | -35.2077 | -31.5438
20.4188 | -35.2081 | -31.5509
20.4249 | -35.2086 | -31.5578
20.431 | -35.209 | -31.5645
20.4371 | -35.2095 | -31.5709
20.4432 | -35.2099 | -31.5777
20.4493 | -35.2104 | -31.5846
20.4554 | -35.2109 | -31.5916
20.4615 | -35.2113 | -31.5984
20.4676 | -35.2117 | -31.6052
20.4737 | -35.2122 | -31.6119
20.4798 | -35.2127 | -31.6188
20.4859 | -35.2132 | -31.6258
20.492 | -35.2137 | -31.6326
20.4981 | -35.2142 | -31.6391
20.5042 | -35.2146 | -31.6462
20.5103 | -35.2151 | -31.6525
20.5164 | -35.2155 | -31.6593
20.5225 | -35.216 | -31.6662
20.5286 | -35.2165 | -31.6731
20.5347 | -35.217 | -31.6799
20.5408 | -35.2174 | -31.6865
20.5469 | -35.2179 | -31.6933
20.553 | -35.2184 | -31.7002
20.5592 | -35.2189 | -31.7068
20.5653 | -35.2193 | -31.7137
20.5714 | -35.2198 | -31.7205
20.5775 | -35.2203 | -31.7272
20.5836 | -35.2207 | -31.7342
20.5897 | -35.2212 | -31.7411
20.5958 | -35.2217 | -31.7484
20.6019 | -35.2222 | -31.755
20.608 | -35.2226 | -31.762
20.6141 | -35.2231 | -31.7687
20.6202 | -35.2235 | -31.7756
20.6263 | -35.224 | -31.7824
20.6324 | -35.2245 | -31.789
20.6385 | -35.225 | -31.7959
20.6446 | -35.2254 | -31.8028
20.6507 | -35.2259 | -31.8095
20.6568 | -35.2264 | -31.8163
20.6629 | -35.2268 | -31.8229
20.669 | -35.2273 | -31.8297
20.6751 | -35.2278 | -31.8364
20.6812 | -35.2283 | -31.843
20.6873 | -35.2287 | -31.8499
20.6934 | -35.2292 | -31.8566
20.6995 | -35.2297 | -31.8633
20.7057 | -35.2301 | -31.8702
20.7118 | -35.2306 | -31.877
20.7179 | -35.2311 | -31.8835
20.724 | -35.2315 | -31.8904
20.7301 | -35.232 | -31.8971
20.7362 | -35.2324 | -31.9037
20.7423 | -35.2329 | -31.9106
20.7484 | -35.2334 | -31.9177
20.7545 | -35.2339 | -31.924
20.7606 | -35.2344 | -31.9306
20.7667 | -35.2348 | -31.9375
20.7728 | -35.2353 | -31.9442
20.7789 | -35.2358 | -31.9509
20.785 | -35.2363 | -31.9576
20.7911 | -35.2368 | -31.9643
20.7972 | -35.2372 | -31.9712
20.8033 | -35.2377 | -31.9781
20.8094 | -35.2382 | -31.9851
20.8155 | -35.2387 | -31.9918
20.8216 | -35.2391 | -31.9986
20.8277 | -35.2396 | -32.005
20.8338 | -35.2401 | -32.0118
20.8399 | -35.2405 | -32.0185
20.8461 | -35.241 | -32.0252
20.8522 | -35.2415 | -32.0321
20.8583 | -35.242 | -32.0389
20.8644 | -35.2424 | -32.0456
20.8705 | -35.2429 | -32.0524
20.8766 | -35.2433 | -32.0591
20.8827 | -35.2438 | -32.066
20.8888 | -35.2442 | -32.0727
20.8949 | -35.2447 | -32.0792
20.901 | -35.2451 | -32.0859
20.9071 | -35.2456 | -32.0925
20.9132 | -35.2461 | -32.099
20.9193 | -35.2466 | -32.106
20.9254 | -35.247 | -32.1125
20.9315 | -35.2476 | -32.1188
20.9376 | -35.248 | -32.1258
20.9437 | -35.2485 | -32.1323
20.9498 | -35.2489 | -32.1391
20.9559 | -35.2494 | -32.146
20.962 | -35.2498 | -32.1529
20.9681 | -35.2503 | -32.1594
20.9742 | -35.2507 | -32.1662
20.9803 | -35.2512 | -32.173
20.9864 | -35.2516 | -32.1797
20.9926 | -35.2521 | -32.1865
20.9987 | -35.2526 | -32.1929
21.0048 | -35.2531 | -32.1996
21.0109 | -35.2536 | -32.2062
21.017 | -35.254 | -32.2128
21.0231 | -35.2546 | -32.2195
21.0292 | -35.255 | -32.2262
21.0353 | -35.2555 | -32.2326
21.0414 | -35.2559 | -32.2391
21.0475 | -35.2564 | -32.2459
21.0536 | -35.2568 | -32.2524
21.0597 | -35.2573 | -32.2591
21.0658 | -35.2578 | -32.2658
21.0719 | -35.2583 | -32.2723
21.078 | -35.2587 | -32.2788
21.0841 | -35.2592 | -32.2856
21.0902 | -35.2596 | -32.2923
21.0963 | -35.2601 | -32.2992
21.1024 | -35.2606 | -32.3058
21.1085 | -35.261 | -32.3124
21.1146 | -35.2615 | -32.3193
21.1207 | -35.262 | -32.3261
21.1268 | -35.2624 | -32.3328
21.133 | -35.2629 | -32.3395
21.1391 | -35.2634 | -32.3458
21.1452 | -35.2638 | -32.3525
21.1513 | -35.2643 | -32.3591
21.1574 | -35.2647 | -32.366
21.1635 | -35.2652 | -32.3728
21.1696 | -35.2657 | -32.3793
21.1757 | -35.2661 | -32.3859
21.1818 | -35.2666 | -32.3927
21.1879 | -35.2671 | -32.3994
21.194 | -35.2675 | -32.4061
21.2001 | -35.268 | -32.4129
21.2062 | -35.2685 | -32.4195
21.2123 | -35.269 | -32.4259
21.2184 | -35.2695 | -32.4325
21.2245 | -35.27 | -32.439
21.2306 | -35.2705 | -32.4456
21.2367 | -35.2709 | -32.4524
21.2428 | -35.2714 | -32.4591
21.2489 | -35.2719 | -32.4657
21.255 | -35.2723 | -32.4723
21.2611 | -35.2728 | -32.4787
21.2672 | -35.2732 | -32.4855
21.2733 | -35.2737 | -32.4922
21.2795 | -35.2741 | -32.4985
21.2856 | -35.2746 | -32.5049
21.2917 | -35.275 | -32.5115
21.2978 | -35.2755 | -32.5183
21.3039 | -35.276 | -32.5251
21.31 | -35.2765 | -32.5318
21.3161 | -35.277 | -32.5383
21.3222 | -35.2775 | -32.5449
21.3283 | -35.2779 | -32.5513
21.3344 | -35.2784 | -32.5581
21.3405 | -35.2789 | -32.565
21.3466 | -35.2793 | -32.5719
21.3527 | -35.2798 | -32.5783
21.3588 | -35.2803 | -32.5849
21.3649 | -35.2807 | -32.5913
21.371 | -35.2812 | -32.5977
21.3771 | -35.2816 | -32.6042
21.3832 | -35.2821 | -32.611
21.3893 | -35.2825 | -32.6174
21.3954 | -35.283 | -32.6242
21.4015 | -35.2835 | -32.6309
21.4076 | -35.284 | -32.6375
21.4137 | -35.2845 | -32.6441
21.4199 | -35.2849 | -32.6507
21.426 | -35.2854 | -32.6571
21.4321 | -35.2859 | -32.6635
21.4382 | -35.2863 | -32.67
21.4443 | -35.2868 | -32.6766
21.4504 | -35.2873 | -32.6832
21.4565 | -35.2877 | -32.6894
21.4626 | -35.2882 | -32.6961
21.4687 | -35.2887 | -32.7027
21.4748 | -35.2891 | -32.7092
21.4809 | -35.2896 | -32.7156
21.487 | -35.2901 | -32.7222
21.4931 | -35.2905 | -32.7287
21.4992 | -35.291 | -32.7353
21.5053 | -35.2915 | -32.7419
21.5114 | -35.2919 | -32.7484
21.5175 | -35.2924 | -32.755
21.5236 | -35.2929 | -32.7613
21.5297 | -35.2934 | -32.7675
21.5358 | -35.2938 | -32.7739
21.5419 | -35.2943 | -32.7807
21.548 | -35.2948 | -32.7879
21.5541 | -35.2952 | -32.7945
21.5602 | -35.2957 | -32.8012
21.5664 | -35.2962 | -32.8076
21.5725 | -35.2966 | -32.8141
21.5786 | -35.2971 | -32.8207
21.5847 | -35.2975 | -32.8272
21.5908 | -35.298 | -32.8335
21.5969 | -35.2985 | -32.8401
21.603 | -35.299 | -32.8469
21.6091 | -35.2994 | -32.8533
21.6152 | -35.2999 | -32.8599
21.6213 | -35.3004 | -32.8664
21.6274 | -35.3009 | -32.8729
21.6335 | -35.3014 | -32.879
21.6396 | -35.3018 | -32.8856
21.6457 | -35.3023 | -32.8921
21.6518 | -35.3027 | -32.8986
21.6579 | -35.3032 | -32.9052
21.664 | -35.3037 | -32.912
21.6701 | -35.3041 | -32.918
21.6762 | -35.3046 | -32.9246
21.6823 | -35.3051 | -32.9312
21.6884 | -35.3056 | -32.9379
21.6945 | -35.306 | -32.9446
21.7006 | -35.3065 | -32.9513
21.7068 | -35.307 | -32.9579
21.7129 | -35.3075 | -32.9643
21.719 | -35.308 | -32.9708
21.7251 | -35.3084 | -32.9775
21.7312 | -35.3089 | -32.9842
21.7373 | -35.3094 | -32.9908
21.7434 | -35.3098 | -32.9974
21.7495 | -35.3103 | -33.0036
21.7556 | -35.3108 | -33.0102
21.7617 | -35.3113 | -33.0171
21.7678 | -35.3117 | -33.0234
21.7739 | -35.3122 | -33.0299
21.78 | -35.3127 | -33.0365
21.7861 | -35.3132 | -33.043
21.7922 | -35.3137 | -33.0495
21.7983 | -35.3141 | -33.0559
21.8044 | -35.3146 | -33.0623
21.8105 | -35.3151 | -33.0685
21.8166 | -35.3156 | -33.0749
21.8227 | -35.3161 | -33.0814
21.8288 | -35.3166 | -33.088
21.8349 | -35.317 | -33.0944
21.841 | -35.3175 | -33.1008
21.8471 | -35.318 | -33.1068
21.8533 | -35.3185 | -33.1136
21.8594 | -35.3189 | -33.12
21.8655 | -35.3194 | -33.1264
21.8716 | -35.3198 | -33.1328
21.8777 | -35.3203 | -33.1391
21.8838 | -35.3208 | -33.1456
21.8899 | -35.3212 | -33.1523
21.896 | -35.3217 | -33.1588
21.9021 | -35.3222 | -33.1652
21.9082 | -35.3226 | -33.1719
21.9143 | -35.3231 | -33.1782
21.9204 | -35.3236 | -33.1846
21.9265 | -35.3241 | -33.1909
21.9326 | -35.3246 | -33.1975
21.9387 | -35.3251 | -33.2039
21.9448 | -35.3255 | -33.2104
21.9509 | -35.326 | -33.2166
21.957 | -35.3265 | -33.2232
21.9631 | -35.327 | -33.2296
21.9692 | -35.3274 | -33.2361
21.9753 | -35.3279 | -33.2426
21.9814 | -35.3284 | -33.2493
21.9875 | -35.3289 | -33.2553
21.9937 | -35.3293 | -33.262
21.9998 | -35.3298 | -33.2687
22.0059 | -35.3303 | -33.2753
22.012 | -35.3308 | -33.2818
22.0181 | -35.3313 | -33.2878
22.0242 | -35.3317 | -33.2944
22.0303 | -35.3321 | -33.301
22.0364 | -35.3326 | -33.3074
22.0425 | -35.3331 | -33.3137
22.0486 | -35.3336 | -33.3202
22.0547 | -35.334 | -33.3267
22.0608 | -35.3345 | -33.3332
22.0669 | -35.335 | -33.3397
22.073 | -35.3354 | -33.346
22.0791 | -35.3359 | -33.3523
22.0852 | -35.3364 | -33.3587
22.0913 | -35.3369 | -33.365
22.0974 | -35.3373 | -33.3717
22.1035 | -35.3377 | -33.3781
22.1096 | -35.3382 | -33.3846
22.1157 | -35.3386 | -33.3915
22.1218 | -35.3391 | -33.3977
22.1279 | -35.3395 | -33.4038
22.134 | -35.3401 | -33.4104
22.1402 | -35.3405 | -33.4167
22.1463 | -35.341 | -33.4231
22.1524 | -35.3415 | -33.4297
22.1585 | -35.342 | -33.4362
22.1646 | -35.3425 | -33.4426
22.1707 | -35.343 | -33.449
22.1768 | -35.3434 | -33.4554
22.1829 | -35.3439 | -33.4618
22.189 | -35.3443 | -33.4684
22.1951 | -35.3448 | -33.4747
22.2012 | -35.3453 | -33.4812
22.2073 | -35.3458 | -33.4876
22.2134 | -35.3462 | -33.4941
22.2195 | -35.3467 | -33.5004
22.2256 | -35.3472 | -33.5069
22.2317 | -35.3476 | -33.5131
22.2378 | -35.3481 | -33.5196
22.2439 | -35.3486 | -33.5261
22.25 | -35.349 | -33.5326
22.2561 | -35.3495 | -33.5389
22.2622 | -35.3499 | -33.5453
22.2683 | -35.3504 | -33.5518
22.2744 | -35.3509 | -33.558
22.2806 | -35.3513 | -33.5644
22.2867 | -35.3518 | -33.5707
22.2928 | -35.3523 | -33.5769
22.2989 | -35.3527 | -33.5833
22.305 | -35.3532 | -33.5897
22.3111 | -35.3537 | -33.5959
22.3172 | -35.3542 | -33.6021
22.3233 | -35.3546 | -33.6086
22.3294 | -35.3551 | -33.6152
22.3355 | -35.3556 | -33.6213
22.3416 | -35.3561 | -33.6276
22.3477 | -35.3566 | -33.634
22.3538 | -35.3571 | -33.6405
22.3599 | -35.3576 | -33.6468
22.366 | -35.358 | -33.6529
22.3721 | -35.3585 | -33.6588
22.3782 | -35.359 | -33.6652
22.3843 | -35.3595 | -33.6717
22.3904 | -35.3599 | -33.678
22.3965 | -35.3604 | -33.6843
22.4026 | -35.3608 | -33.6909
22.4087 | -35.3613 | -33.6971
22.4148 | -35.3618 | -33.7035
22.4209 | -35.3623 | -33.7098
22.4271 | -35.3627 | -33.7161
22.4332 | -35.3632 | -33.7223
22.4393 | -35.3637 | -33.7285
22.4454 | -35.3642 | -33.7347
22.4515 | -35.3647 | -33.741
22.4576 | -35.3652 | -33.7474
22.4637 | -35.3656 | -33.7538
22.4698 | -35.3661 | -33.7601
22.4759 | -35.3666 | -33.7665
22.482 | -35.3671 | -33.7731
22.4881 | -35.3675 | -33.7795
22.4942 | -35.368 | -33.7859
22.5003 | -35.3685 | -33.7922
22.5064 | -35.3689 | -33.7985
22.5125 | -35.3694 | -33.8046
22.5186 | -35.3699 | -33.8108
22.5247 | -35.3703 | -33.8172
22.5308 | -35.3708 | -33.8236
22.5369 | -35.3712 | -33.8301
22.543 | -35.3717 | -33.8366
22.5491 | -35.3722 | -33.8426
22.5552 | -35.3726 | -33.849
22.5613 | -35.3731 | -33.8554
22.5675 | -35.3736 | -33.8616
22.5736 | -35.3741 | -33.8681
22.5797 | -35.3745 | -33.8744
22.5858 | -35.375 | -33.8805
22.5919 | -35.3755 | -33.8868
22.598 | -35.376 | -33.8934
22.6041 | -35.3764 | -33.8998
22.6102 | -35.3769 | -33.906
22.6163 | -35.3774 | -33.9121
22.6224 | -35.3778 | -33.9185
22.6285 | -35.3783 | -33.9248
22.6346 | -35.3788 | -33.9311
22.6407 | -35.3792 | -33.9373
22.6468 | -35.3797 | -33.9436
22.6529 | -35.3802 | -33.9497
22.659 | -35.3807 | -33.9559
22.6651 | -35.3811 | -33.9622
22.6712 | -35.3816 | -33.9684
22.6773 | -35.3821 | -33.9748
22.6834 | -35.3826 | -33.9811
22.6895 | -35.383 | -33.9871
22.6956 | -35.3835 | -33.9932
22.7017 | -35.384 | -33.9995
22.7079 | -35.3844 | -34.0057
22.714 | -35.3849 | -34.0118
22.7201 | -35.3854 | -34.0183
22.7262 | -35.3859 | -34.0244
22.7323 | -35.3863 | -34.0306
22.7384 | -35.3868 | -34.037
22.7445 | -35.3873 | -34.0434
22.7506 | -35.3877 | -34.0496
22.7567 | -35.3882 | -34.0558
22.7628 | -35.3887 | -34.062
22.7689 | -35.3892 | -34.0682
22.775 | -35.3897 | -34.0744
22.7811 | -35.3902 | -34.0807
22.7872 | -35.3906 | -34.087
22.7933 | -35.3911 | -34.0931
22.7994 | -35.3915 | -34.0995
22.8055 | -35.392 | -34.1057
22.8116 | -35.3925 | -34.1118
22.8177 | -35.393 | -34.1182
22.8238 | -35.3935 | -34.1244
22.8299 | -35.3939 | -34.1304
22.836 | -35.3944 | -34.1366
22.8421 | -35.3949 | -34.143
22.8482 | -35.3953 | -34.1494
22.8544 | -35.3958 | -34.1556
22.8605 | -35.3962 | -34.1619
22.8666 | -35.3967 | -34.1681
22.8727 | -35.3972 | -34.1744
22.8788 | -35.3976 | -34.1807
22.8849 | -35.3981 | -34.1869
22.891 | -35.3985 | -34.1933
22.8971 | -35.399 | -34.1996
22.9032 | -35.3995 | -34.2056
22.9093 | -35.3999 | -34.2117
22.9154 | -35.4004 | -34.2181
22.9215 | -35.4009 | -34.2242
22.9276 | -35.4013 | -34.2305
22.9337 | -35.4018 | -34.2363
22.9398 | -35.4023 | -34.2424
22.9459 | -35.4028 | -34.2488
22.952 | -35.4032 | -34.255
22.9581 | -35.4037 | -34.2613
22.9642 | -35.4042 | -34.2676
22.9703 | -35.4046 | -34.2737
22.9764 | -35.4051 | -34.2799
22.9825 | -35.4056 | -34.286
22.9886 | -35.4061 | -34.2924
22.9948 | -35.4065 | -34.2985
23.0009 | -35.407 | -34.3048
23.007 | -35.4075 | -34.3109
23.0131 | -35.4079 | -34.3171
23.0192 | -35.4084 | -34.3233
23.0253 | -35.4088 | -34.3296
23.0314 | -35.4093 | -34.3358
23.0375 | -35.4098 | -34.3421
23.0436 | -35.4102 | -34.3482
23.0497 | -35.4107 | -34.3545
23.0558 | -35.4112 | -34.3607
23.0619 | -35.4116 | -34.3669
23.068 | -35.4121 | -34.373
23.0741 | -35.4125 | -34.379
23.0802 | -35.413 | -34.3853
23.0863 | -35.4134 | -34.3914
23.0924 | -35.4139 | -34.3975
23.0985 | -35.4143 | -34.4036
23.1046 | -35.4148 | -34.4098
23.1107 | -35.4153 | -34.4159
23.1168 | -35.4157 | -34.4222
23.1229 | -35.4162 | -34.4283
23.129 | -35.4167 | -34.4342
23.1351 | -35.4171 | -34.4404
23.1413 | -35.4175 | -34.4464
23.1474 | -35.418 | -34.4525
23.1535 | -35.4184 | -34.4587
23.1596 | -35.4189 | -34.4648
23.1657 | -35.4194 | -34.4709
23.1718 | -35.4199 | -34.4768
23.1779 | -35.4203 | -34.4832
23.184 | -35.4208 | -34.4891
23.1901 | -35.4212 | -34.4952
23.1962 | -35.4217 | -34.5014
23.2023 | -35.4222 | -34.5076
23.2084 | -35.4226 | -34.5138
23.2145 | -35.4231 | -34.5198
23.2206 | -35.4235 | -34.5258
23.2267 | -35.424 | -34.5321
23.2328 | -35.4245 | -34.5382
23.2389 | -35.4249 | -34.5444
23.245 | -35.4254 | -34.5506
23.2511 | -35.4258 | -34.5565
23.2572 | -35.4263 | -34.5626
23.2633 | -35.4268 | -34.5689
23.2694 | -35.4272 | -34.5752
23.2755 | -35.4277 | -34.5813
23.2817 | -35.4282 | -34.5875
23.2878 | -35.4286 | -34.5937
23.2939 | -35.4291 | -34.5997
23.3 | -35.4295 | -34.606
23.3061 | -35.43 | -34.6122
23.3122 | -35.4305 | -34.6184
23.3183 | -35.4309 | -34.6244
23.3244 | -35.4313 | -34.6304
23.3305 | -35.4318 | -34.6367
23.3366 | -35.4323 | -34.6429
23.3427 | -35.4327 | -34.6491
23.3488 | -35.4332 | -34.6554
23.3549 | -35.4337 | -34.6616
23.361 | -35.4341 | -34.6677
23.3671 | -35.4346 | -34.6739
23.3732 | -35.435 | -34.68
23.3793 | -35.4355 | -34.686
23.3854 | -35.4359 | -34.6923
23.3915 | -35.4364 | -34.6982
23.3976 | -35.4369 | -34.7042
23.4037 | -35.4374 | -34.7104
23.4098 | -35.4378 | -34.7165
23.4159 | -35.4383 | -34.7228
23.422 | -35.4388 | -34.729
23.4282 | -35.4392 | -34.7348
23.4343 | -35.4397 | -34.7408
23.4404 | -35.4401 | -34.7469
23.4465 | -35.4406 | -34.753
23.4526 | -35.441 | -34.7591
23.4587 | -35.4415 | -34.7653
23.4648 | -35.442 | -34.7713
23.4709 | -35.4425 | -34.7774
23.477 | -35.4429 | -34.7834
23.4831 | -35.4434 | -34.7894
23.4892 | -35.4439 | -34.7956
23.4953 | -35.4443 | -34.8018
23.5014 | -35.4448 | -34.8079
23.5075 | -35.4452 | -34.814
23.5136 | -35.4457 | -34.8201
23.5197 | -35.4461 | -34.8263
23.5258 | -35.4466 | -34.8323
23.5319 | -35.447 | -34.8381
23.538 | -35.4474 | -34.8442
23.5441 | -35.4479 | -34.8503
23.5502 | -35.4483 | -34.8565
23.5563 | -35.4488 | -34.8627
23.5624 | -35.4493 | -34.8689
23.5686 | -35.4497 | -34.8746
23.5747 | -35.4502 | -34.8807
23.5808 | -35.4506 | -34.8867
23.5869 | -35.4511 | -34.8927
23.593 | -35.4515 | -34.8988
23.5991 | -35.452 | -34.9048
23.6052 | -35.4524 | -34.9109
23.6113 | -35.4529 | -34.917
23.6174 | -35.4533 | -34.9232
23.6235 | -35.4537 | -34.9292
23.6296 | -35.4542 | -34.9353
23.6357 | -35.4546 | -34.9413
23.6418 | -35.4551 | -34.9473
23.6479 | -35.4556 | -34.9534
23.654 | -35.456 | -34.9594
23.6601 | -35.4565 | -34.9655
23.6662 | -35.4569 | -34.9714
23.6723 | -35.4574 | -34.9772
23.6784 | -35.4578 | -34.9832
23.6845 | -35.4583 | -34.9893
23.6906 | -35.4587 | -34.9954
23.6967 | -35.4592 | -35.0013
23.7028 | -35.4596 | -35.0074
23.7089 | -35.4601 | -35.0133
23.7151 | -35.4605 | -35.0195
23.7212 | -35.461 | -35.0255
23.7273 | -35.4614 | -35.0317
23.7334 | -35.4619 | -35.0378
23.7395 | -35.4623 | -35.0438
23.7456 | -35.4628 | -35.0495
23.7517 | -35.4632 | -35.0557
23.7578 | -35.4636 | -35.0618
23.7639 | -35.4641 | -35.0679
23.77 | -35.4645 | -35.0739
23.7761 | -35.4649 | -35.0801
23.7822 | -35.4654 | -35.0859
23.7883 | -35.4658 | -35.092
23.7944 | -35.4663 | -35.098
23.8005 | -35.4667 | -35.1041
23.8066 | -35.4671 | -35.1103
23.8127 | -35.4676 | -35.1159
23.8188 | -35.468 | -35.1219
23.8249 | -35.4685 | -35.128
23.831 | -35.4689 | -35.1341
23.8371 | -35.4694 | -35.1399
23.8432 | -35.4698 | -35.1459
23.8493 | -35.4703 | -35.1517
23.8555 | -35.4707 | -35.1577
23.8616 | -35.4711 | -35.1638
23.8677 | -35.4716 | -35.17
23.8738 | -35.472 | -35.1759
23.8799 | -35.4725 | -35.182
23.886 | -35.4729 | -35.1878
23.8921 | -35.4733 | -35.1937
23.8982 | -35.4738 | -35.1999
23.9043 | -35.4742 | -35.2062
23.9104 | -35.4746 | -35.2122
23.9165 | -35.475 | -35.2181
23.9226 | -35.4755 | -35.224
23.9287 | -35.4759 | -35.2301
23.9348 | -35.4764 | -35.2362
23.9409 | -35.4768 | -35.2423
23.947 | -35.4773 | -35.2483
23.9531 | -35.4777 | -35.2541
23.9592 | -35.4781 | -35.2601
23.9653 | -35.4786 | -35.266
23.9714 | -35.4791 | -35.2719
23.9775 | -35.4795 | -35.2779
23.9836 | -35.4799 | -35.284
23.9897 | -35.4804 | -35.2899
23.9958 | -35.4808 | -35.2958
24.002 | -35.4812 | -35.3018
24.0081 | -35.4817 | -35.3077
24.0142 | -35.4822 | -35.3136
24.0203 | -35.4826 | -35.3196
24.0264 | -35.4831 | -35.3255
24.0325 | -35.4835 | -35.3315
24.0386 | -35.484 | -35.3377
24.0447 | -35.4844 | -35.3437
24.0508 | -35.4848 | -35.3497
24.0569 | -35.4853 | -35.3558
24.063 | -35.4857 | -35.3616
24.0691 | -35.4862 | -35.3676
24.0752 | -35.4866 | -35.3738
24.0813 | -35.4871 | -35.3796
24.0874 | -35.4875 | -35.3855
24.0935 | -35.4879 | -35.3918
24.0996 | -35.4883 | -35.3975
24.1057 | -35.4888 | -35.4036
24.1118 | -35.4892 | -35.4098
24.1179 | -35.4897 | -35.4159
24.124 | -35.4901 | -35.422
24.1301 | -35.4905 | -35.4278
24.1362 | -35.491 | -35.4338
24.1424 | -35.4914 | -35.4398
24.1485 | -35.4918 | -35.4459
24.1546 | -35.4923 | -35.4518
24.1607 | -35.4927 | -35.4578
24.1668 | -35.4931 | -35.4635
24.1729 | -35.4935 | -35.4695
24.179 | -35.494 | -35.4753
24.1851 | -35.4944 | -35.4811
24.1912 | -35.4948 | -35.487
24.1973 | -35.4953 | -35.493
24.2034 | -35.4957 | -35.4988
24.2095 | -35.4961 | -35.5049
24.2156 | -35.4966 | -35.5109
24.2217 | -35.497 | -35.5169
24.2278 | -35.4974 | -35.5229
24.2339 | -35.4978 | -35.5288
24.24 | -35.4983 | -35.5344
24.2461 | -35.4987 | -35.5404
24.2522 | -35.4991 | -35.5464
24.2583 | -35.4996 | -35.5523
24.2644 | -35.5 | -35.5582
24.2705 | -35.5004 | -35.564
24.2766 | -35.5008 | -35.5698
24.2827 | -35.5012 | -35.5759
24.2889 | -35.5017 | -35.582
24.295 | -35.502 | -35.5879
24.3011 | -35.5025 | -35.5938
24.3072 | -35.5029 | -35.5998
24.3133 | -35.5033 | -35.6056
24.3194 | -35.5038 | -35.6116
24.3255 | -35.5042 | -35.6177
24.3316 | -35.5046 | -35.6235
24.3377 | -35.505 | -35.6292
24.3438 | -35.5054 | -35.6349
24.3499 | -35.5058 | -35.6407
24.356 | -35.5062 | -35.6466
24.3621 | -35.5067 | -35.6525
24.3682 | -35.5071 | -35.6584
24.3743 | -35.5075 | -35.6644
24.3804 | -35.5079 | -35.6701
24.3865 | -35.5084 | -35.6761
24.3926 | -35.5088 | -35.6821
24.3987 | -35.5092 | -35.6881
24.4048 | -35.5097 | -35.6941
24.4109 | -35.51 | -35.7
24.417 | -35.5105 | -35.7058
24.4231 | -35.5109 | -35.7119
24.4293 | -35.5114 | -35.7177
24.4354 | -35.5118 | -35.7235
24.4415 | -35.5122 | -35.7294
24.4476 | -35.5126 | -35.7348
24.4537 | -35.513 | -35.7406
24.4598 | -35.5135 | -35.7464
24.4659 | -35.5139 | -35.7524
24.472 | -35.5143 | -35.7582
24.4781 | -35.5147 | -35.7642
24.4842 | -35.5151 | -35.7699
24.4903 | -35.5156 | -35.7758
24.4964 | -35.516 | -35.7817
24.5025 | -35.5164 | -35.7874
24.5086 | -35.5168 | -35.7931
24.5147 | -35.5173 | -35.7989
24.5208 | -35.5177 | -35.8044
24.5269 | -35.5181 | -35.8102
24.533 | -35.5185 | -35.8161
24.5391 | -35.5189 | -35.8219
24.5452 | -35.5194 | -35.8277
24.5513 | -35.5198 | -35.8334
24.5574 | -35.5202 | -35.8392
24.5635 | -35.5207 | -35.8453
24.5696 | -35.5211 | -35.8511
24.5758 | -35.5215 | -35.857
24.5819 | -35.522 | -35.863
24.588 | -35.5223 | -35.8688
24.5941 | -35.5228 | -35.8745
24.6002 | -35.5232 | -35.8804
24.6063 | -35.5237 | -35.8864
24.6124 | -35.5241 | -35.8922
24.6185 | -35.5245 | -35.8981
24.6246 | -35.5249 | -35.9038
24.6307 | -35.5253 | -35.9097
24.6368 | -35.5258 | -35.9157
24.6429 | -35.5262 | -35.9214
24.649 | -35.5266 | -35.9271
24.6551 | -35.5271 | -35.933
24.6612 | -35.5275 | -35.9386
24.6673 | -35.528 | -35.9443
24.6734 | -35.5284 | -35.9503
24.6795 | -35.5288 | -35.9562
24.6856 | -35.5292 | -35.9621
24.6917 | -35.5296 | -35.9679
24.6978 | -35.5301 | -35.9734
24.7039 | -35.5305 | -35.9792
24.71 | -35.5309 | -35.9851
24.7162 | -35.5313 | -35.991
24.7223 | -35.5318 | -35.9969
24.7284 | -35.5322 | -36.0025
24.7345 | -35.5326 | -36.0082
24.7406 | -35.533 | -36.0141
24.7467 | -35.5334 | -36.0199
24.7528 | -35.5338 | -36.0257
24.7589 | -35.5343 | -36.0315
24.765 | -35.5347 | -36.037
24.7711 | -35.5351 | -36.0428
24.7772 | -35.5356 | -36.0487
24.7833 | -35.536 | -36.0543
24.7894 | -35.5364 | -36.0601
24.7955 | -35.5368 | -36.0659
24.8016 | -35.5372 | -36.0714
24.8077 | -35.5376 | -36.077
24.8138 | -35.5381 | -36.083
24.8199 | -35.5385 | -36.0888
24.826 | -35.5389 | -36.0945
24.8321 | -35.5393 | -36.1005
24.8382 | -35.5397 | -36.1059
24.8443 | -35.5401 | -36.1117
24.8504 | -35.5405 | -36.1175
24.8565 | -35.541 | -36.1233
24.8627 | -35.5414 | -36.1291
24.8688 | -35.5418 | -36.1347
24.8749 | -35.5422 | -36.1405
24.881 | -35.5426 | -36.1464
24.8871 | -35.543 | -36.1522
24.8932 | -35.5434 | -36.1581
24.8993 | -35.5439 | -36.1638
24.9054 | -35.5443 | -36.1694
24.9115 | -35.5447 | -36.1751
24.9176 | -35.5451 | -36.1809
24.9237 | -35.5455 | -36.1867
24.9298 | -35.5459 | -36.1924
24.9359 | -35.5463 | -36.1984
24.942 | -35.5467 | -36.204
24.9481 | -35.5471 | -36.2099
24.9542 | -35.5476 | -36.2155
24.9603 | -35.5479 | -36.2213
24.9664 | -35.5484 | -36.2271
24.9725 | -35.5488 | -36.233
24.9786 | -35.5492 | -36.2385
24.9847 | -35.5496 | -36.2443
24.9908 | -35.55 | -36.2501
24.9969 | -35.5504 | -36.256
25.0031 | -35.5508 | -36.2617
25.0092 | -35.5512 | -36.2673
25.0153 | -35.5516 | -36.273
25.0214 | -35.552 | -36.2788
25.0275 | -35.5524 | -36.2846
25.0336 | -35.5529 | -36.2903
25.0397 | -35.5533 | -36.2961
25.0458 | -35.5537 | -36.3017
25.0519 | -35.5541 | -36.3075
25.058 | -35.5545 | -36.3134
25.0641 | -35.5549 | -36.3192
25.0702 | -35.5553 | -36.325
25.0763 | -35.5557 | -36.3307
25.0824 | -35.5561 | -36.3361
25.0885 | -35.5565 | -36.3416
25.0946 | -35.5569 | -36.3474
25.1007 | -35.5573 | -36.3531
25.1068 | -35.5578 | -36.3588
25.1129 | -35.5582 | -36.3647
25.119 | -35.5586 | -36.3702
25.1251 | -35.559 | -36.376
25.1312 | -35.5594 | -36.3817
25.1373 | -35.5598 | -36.3874
25.1435 | -35.5602 | -36.3933
25.1496 | -35.5606 | -36.3987
25.1557 | -35.5611 | -36.4045
25.1618 | -35.5615 | -36.4102
25.1679 | -35.5619 | -36.416
25.174 | -35.5623 | -36.4219
25.1801 | -35.5627 | -36.4275
25.1862 | -35.5631 | -36.433
25.1923 | -35.5635 | -36.4388
25.1984 | -35.5639 | -36.4445
25.2045 | -35.5643 | -36.4502
25.2106 | -35.5647 | -36.4559
25.2167 | -35.5651 | -36.4618
25.2228 | -35.5655 | -36.4676
25.2289 | -35.5659 | -36.4733
25.235 | -35.5663 | -36.4791
25.2411 | -35.5667 | -36.485
25.2472 | -35.5671 | -36.4907
25.2533 | -35.5675 | -36.4965
25.2594 | -35.5679 | -36.5018
25.2655 | -35.5683 | -36.5075
25.2716 | -35.5687 | -36.5132
25.2777 | -35.5691 | -36.519
25.2838 | -35.5695 | -36.5248
25.29 | -35.5699 | -36.5303
25.2961 | -35.5703 | -36.5357
25.3022 | -35.5707 | -36.5413
25.3083 | -35.5711 | -36.547
25.3144 | -35.5715 | -36.5528
25.3205 | -35.5719 | -36.5585
25.3266 | -35.5723 | -36.564
25.3327 | -35.5727 | -36.5696
25.3388 | -35.573 | -36.5754
25.3449 | -35.5734 | -36.5811
25.351 | -35.5738 | -36.5867
25.3571 | -35.5743 | -36.5923
25.3632 | -35.5746 | -36.5977
25.3693 | -35.575 | -36.6033
25.3754 | -35.5754 | -36.6091
25.3815 | -35.5758 | -36.6147
25.3876 | -35.5762 | -36.6204
25.3937 | -35.5766 | -36.6261
25.3998 | -35.577 | -36.6315
25.4059 | -35.5774 | -36.6371
25.412 | -35.5778 | -36.6426
25.4181 | -35.5782 | -36.6483
25.4242 | -35.5786 | -36.6541
25.4304 | -35.579 | -36.6597
25.4365 | -35.5793 | -36.6652
25.4426 | -35.5797 | -36.6709
25.4487 | -35.5801 | -36.6766
25.4548 | -35.5805 | -36.6821
25.4609 | -35.5809 | -36.6877
25.467 | -35.5813 | -36.6932
25.4731 | -35.5817 | -36.6988
25.4792 | -35.5821 | -36.7046
25.4853 | -35.5825 | -36.7102
25.4914 | -35.5829 | -36.716
25.4975 | -35.5833 | -36.7215
25.5036 | -35.5837 | -36.7269
25.5097 | -35.5841 | -36.7326
25.5158 | -35.5845 | -36.7382
25.5219 | -35.5849 | -36.744
25.528 | -35.5853 | -36.7495
25.5341 | -35.5857 | -36.7552
25.5402 | -35.5861 | -36.7607
25.5463 | -35.5864 | -36.7664
25.5524 | -35.5868 | -36.7721
25.5585 | -35.5872 | -36.7778
25.5646 | -35.5876 | -36.7835
25.5707 | -35.588 | -36.789
25.5769 | -35.5884 | -36.7945
25.583 | -35.5888 | -36.8001
25.5891 | -35.5892 | -36.8058
25.5952 | -35.5895 | -36.8115
25.6013 | -35.5899 | -36.8171
25.6074 | -35.5903 | -36.8224
25.6135 | -35.5907 | -36.8281
25.6196 | -35.5911 | -36.8339
25.6257 | -35.5915 | -36.8394
25.6318 | -35.5919 | -36.8448
25.6379 | -35.5923 | -36.8506
25.644 | -35.5927 | -36.856
25.6501 | -35.593 | -36.8616
25.6562 | -35.5934 | -36.8671
25.6623 | -35.5938 | -36.8726
25.6684 | -35.5942 | -36.8782
25.6745 | -35.5946 | -36.884
25.6806 | -35.595 | -36.8893
25.6867 | -35.5955 | -36.8944
25.6928 | -35.5958 | -36.8988
25.6989 | -35.5962 | -36.9051
25.705 | -35.5968 | -36.911
25.7111 | -35.5975 | -36.9178
25.7173 | -35.5977 | -36.9227
25.7234 | -35.598 | -36.9282
25.7295 | -35.5984 | -36.9338
25.7356 | -35.5988 | -36.9394
25.7417 | -35.5992 | -36.9451
25.7478 | -35.5996 | -36.9505
25.7539 | -35.6 | -36.9558
25.76 | -35.6004 | -36.9615
25.7661 | -35.6008 | -36.9669
25.7722 | -35.6012 | -36.9724
25.7783 | -35.6016 | -36.978
25.7844 | -35.602 | -36.9834
25.7905 | -35.6023 | -36.9889
25.7966 | -35.6027 | -36.9945
25.8027 | -35.6031 | -37.0001
25.8088 | -35.6035 | -37.0059
25.8149 | -35.6039 | -37.0114
25.821 | -35.6043 | -37.017
25.8271 | -35.6047 | -37.0224
25.8332 | -35.6051 | -37.028
25.8393 | -35.6055 | -37.0336
25.8454 | -35.6059 | -37.039
25.8515 | -35.6062 | -37.0445
25.8576 | -35.6066 | -37.0498
25.8638 | -35.607 | -37.0553
25.8699 | -35.6074 | -37.0609
25.876 | -35.6078 | -37.0662
25.8821 | -35.6081 | -37.0718
25.8882 | -35.6085 | -37.077
25.8943 | -35.6089 | -37.0826
25.9004 | -35.6093 | -37.0883
25.9065 | -35.6097 | -37.0938
25.9126 | -35.6101 | -37.0993
25.9187 | -35.6105 | -37.1047
25.9248 | -35.6109 | -37.11
25.9309 | -35.6112 | -37.1155
25.937 | -35.6116 | -37.1211
25.9431 | -35.612 | -37.1266
25.9492 | -35.6124 | -37.1321
25.9553 | -35.6128 | -37.1377
25.9614 | -35.6132 | -37.143
25.9675 | -35.6136 | -37.1486
25.9736 | -35.614 | -37.1543
25.9797 | -35.6144 | -37.1599
25.9858 | -35.6147 | -37.1654
25.9919 | -35.6151 | -37.1712
25.998 | -35.6155 | -37.1765
26.0042 | -35.6159 | -37.1821
26.0103 | -35.6163 | -37.1876
26.0164 | -35.6167 | -37.1932
26.0225 | -35.6171 | -37.1986
26.0286 | -35.6175 | -37.2041
26.0347 | -35.6178 | -37.2093
26.0408 | -35.6182 | -37.215
26.0469 | -35.6186 | -37.2204
26.053 | -35.619 | -37.226
26.0591 | -35.6194 | -37.2315
26.0652 | -35.6197 | -37.2368
26.0713 | -35.6201 | -37.2422
26.0774 | -35.6205 | -37.2476
26.0835 | -35.6209 | -37.2532
26.0896 | -35.6213 | -37.2586
26.0957 | -35.6217 | -37.2641
26.1018 | -35.6221 | -37.2694
26.1079 | -35.6224 | -37.275
26.114 | -35.6228 | -37.2804
26.1201 | -35.6232 | -37.2859
26.1262 | -35.6236 | -37.2913
26.1323 | -35.6239 | -37.2968
26.1384 | -35.6243 | -37.3022
26.1445 | -35.6247 | -37.3078
26.1507 | -35.6251 | -37.3134
26.1568 | -35.6255 | -37.3188
26.1629 | -35.6259 | -37.3243
26.169 | -35.6263 | -37.3298
26.1751 | -35.6266 | -37.3351
26.1812 | -35.627 | -37.3405
26.1873 | -35.6274 | -37.346
26.1934 | -35.6278 | -37.3515
26.1995 | -35.6282 | -37.357
26.2056 | -35.6285 | -37.3624
26.2117 | -35.6289 | -37.3678
26.2178 | -35.6293 | -37.3733
26.2239 | -35.6297 | -37.3788
26.23 | -35.63 | -37.3841
26.2361 | -35.6304 | -37.3896
26.2422 | -35.6308 | -37.395
26.2483 | -35.6312 | -37.4004
26.2544 | -35.6315 | -37.4057
26.2605 | -35.6319 | -37.4112
26.2666 | -35.6323 | -37.4167
26.2727 | -35.6327 | -37.4222
26.2788 | -35.6331 | -37.4273
26.2849 | -35.6335 | -37.4326
26.2911 | -35.6338 | -37.438
26.2972 | -35.6342 | -37.4435
26.3033 | -35.6346 | -37.4489
26.3094 | -35.635 | -37.4545
26.3155 | -35.6354 | -37.4598
26.3216 | -35.6357 | -37.4653
26.3277 | -35.6361 | -37.471
26.3338 | -35.6365 | -37.4763
26.3399 | -35.6369 | -37.4816
26.346 | -35.6373 | -37.4868
26.3521 | -35.6376 | -37.4923
26.3582 | -35.638 | -37.4977
26.3643 | -35.6384 | -37.5031
26.3704 | -35.6388 | -37.5087
26.3765 | -35.6392 | -37.514
26.3826 | -35.6396 | -37.5192
26.3887 | -35.6399 | -37.5247
26.3948 | -35.6403 | -37.53
26.4009 | -35.6407 | -37.5354
26.407 | -35.6411 | -37.5408
26.4131 | -35.6415 | -37.5462
26.4192 | -35.6419 | -37.5514
26.4253 | -35.6423 | -37.557
26.4314 | -35.6426 | -37.5623
26.4376 | -35.643 | -37.5678
26.4437 | -35.6434 | -37.5733
26.4498 | -35.6438 | -37.5788
26.4559 | -35.6441 | -37.584
26.462 | -35.6445 | -37.5896
26.4681 | -35.6449 | -37.5949
26.4742 | -35.6453 | -37.6003
26.4803 | -35.6456 | -37.606
26.4864 | -35.646 | -37.6112
26.4925 | -35.6464 | -37.6165
26.4986 | -35.6468 | -37.6219
26.5047 | -35.6471 | -37.6272
26.5108 | -35.6475 | -37.6325
26.5169 | -35.6479 | -37.638
26.523 | -35.6482 | -37.6432
26.5291 | -35.6486 | -37.6486
26.5352 | -35.649 | -37.6539
26.5413 | -35.6493 | -37.6594
26.5474 | -35.6497 | -37.6648
26.5535 | -35.6501 | -37.6702
26.5596 | -35.6504 | -37.6754
26.5657 | -35.6508 | -37.6809
26.5718 | -35.6512 | -37.6862
26.578 | -35.6516 | -37.6917
26.5841 | -35.652 | -37.697
26.5902 | -35.6523 | -37.7024
26.5963 | -35.6527 | -37.7077
26.6024 | -35.6531 | -37.7133
26.6085 | -35.6534 | -37.7188
26.6146 | -35.6538 | -37.7242
26.6207 | -35.6542 | -37.7297
26.6268 | -35.6546 | -37.7352
26.6329 | -35.6549 | -37.7406
26.639 | -35.6553 | -37.7461
26.6451 | -35.6557 | -37.7515
26.6512 | -35.656 | -37.7568
26.6573 | -35.6564 | -37.7623
26.6634 | -35.6568 | -37.7675
26.6695 | -35.6571 | -37.7729
26.6756 | -35.6575 | -37.7785
26.6817 | -35.6578 | -37.7841
26.6878 | -35.6582 | -37.7896
26.6939 | -35.6586 | -37.795
26.7 | -35.6589 | -37.8003
26.7061 | -35.6593 | -37.8056
26.7122 | -35.6597 | -37.8109
26.7183 | -35.6601 | -37.8161
26.7245 | -35.6604 | -37.8215
26.7306 | -35.6608 | -37.8269
26.7367 | -35.6611 | -37.8321
26.7428 | -35.6615 | -37.8374
26.7489 | -35.6619 | -37.8427
26.755 | -35.6622 | -37.848
26.7611 | -35.6626 | -37.8534
26.7672 | -35.663 | -37.8587
26.7733 | -35.6633 | -37.864
26.7794 | -35.6637 | -37.8694
26.7855 | -35.6641 | -37.8747
26.7916 | -35.6644 | -37.88
26.7977 | -35.6648 | -37.8854
26.8038 | -35.6652 | -37.8907
26.8099 | -35.6655 | -37.8961
26.816 | -35.6659 | -37.9016
26.8221 | -35.6662 | -37.9071
26.8282 | -35.6666 | -37.9124
26.8343 | -35.667 | -37.9176
26.8404 | -35.6673 | -37.9229
26.8465 | -35.6677 | -37.9283
26.8526 | -35.6681 | -37.9338
26.8587 | -35.6684 | -37.9391
26.8649 | -35.6688 | -37.9444
26.871 | -35.6692 | -37.9497
26.8771 | -35.6695 | -37.955
26.8832 | -35.6699 | -37.9604
26.8893 | -35.6703 | -37.9658
26.8954 | -35.6707 | -37.9711
26.9015 | -35.671 | -37.9764
26.9076 | -35.6714 | -37.9817
26.9137 | -35.6718 | -37.987
26.9198 | -35.6721 | -37.9925
26.9259 | -35.6725 | -37.998
26.932 | -35.6728 | -38.0033
26.9381 | -35.6732 | -38.0086
26.9442 | -35.6735 | -38.0139
26.9503 | -35.6739 | -38.0192
26.9564 | -35.6743 | -38.0246
26.9625 | -35.6746 | -38.0299
26.9686 | -35.675 | -38.0352
26.9747 | -35.6754 | -38.0405
26.9808 | -35.6757 | -38.0456
26.9869 | -35.6761 | -38.051
26.993 | -35.6764 | -38.0563
26.9991 | -35.6768 | -38.0618
27.0052 | -35.6771 | -38.0671
27.0114 | -35.6775 | -38.0726
27.0175 | -35.6778 | -38.078
27.0236 | -35.6782 | -38.0833
27.0297 | -35.6785 | -38.0887
27.0358 | -35.6789 | -38.094
27.0419 | -35.6793 | -38.0994
27.048 | -35.6796 | -38.1048
27.0541 | -35.68 | -38.11
27.0602 | -35.6803 | -38.1153
27.0663 | -35.6807 | -38.1206
27.0724 | -35.6811 | -38.1258
27.0785 | -35.6814 | -38.1311
27.0846 | -35.6818 | -38.1364
27.0907 | -35.6822 | -38.1418
27.0968 | -35.6825 | -38.147
27.1029 | -35.6829 | -38.1524
27.109 | -35.6832 | -38.158
27.1151 | -35.6836 | -38.1632
27.1212 | -35.6839 | -38.1683
27.1273 | -35.6843 | -38.1738
27.1334 | -35.6847 | -38.1791
27.1395 | -35.6851 | -38.1845
27.1456 | -35.6854 | -38.1899
27.1518 | -35.6858 | -38.1953
27.1579 | -35.6861 | -38.2006
27.164 | -35.6865 | -38.2058
27.1701 | -35.6869 | -38.2111
27.1762 | -35.6872 | -38.2164
27.1823 | -35.6876 | -38.2216
27.1884 | -35.6879 | -38.2271
27.1945 | -35.6883 | -38.2324
27.2006 | -35.6887 | -38.2377
27.2067 | -35.689 | -38.2432
27.2128 | -35.6894 | -38.2485
27.2189 | -35.6897 | -38.2537
27.225 | -35.6901 | -38.2591
27.2311 | -35.6904 | -38.2644
27.2372 | -35.6908 | -38.2696
27.2433 | -35.6912 | -38.2749
27.2494 | -35.6915 | -38.2803
27.2555 | -35.6919 | -38.2857
27.2616 | -35.6922 | -38.2908
27.2677 | -35.6926 | -38.2962
27.2738 | -35.693 | -38.3017
27.2799 | -35.6933 | -38.307
27.286 | -35.6937 | -38.3123
27.2921 | -35.694 | -38.3176
27.2983 | -35.6944 | -38.3228
27.3044 | -35.6947 | -38.3283
27.3105 | -35.6951 | -38.3336
27.3166 | -35.6954 | -38.3389
27.3227 | -35.6958 | -38.3443
27.3288 | -35.6961 | -38.3495
27.3349 | -35.6965 | -38.3546
27.341 | -35.6968 | -38.3599
27.3471 | -35.6972 | -38.3651
27.3532 | -35.6976 | -38.3704
27.3593 | -35.6979 | -38.3759
27.3654 | -35.6983 | -38.3814
27.3715 | -35.6986 | -38.3866
27.3776 | -35.699 | -38.392
27.3837 | -35.6994 | -38.3971
27.3898 | -35.6997 | -38.4023
27.3959 | -35.7001 | -38.4076
27.402 | -35.7004 | -38.4127
27.4081 | -35.7008 | -38.418
27.4142 | -35.7012 | -38.4234
27.4203 | -35.7015 | -38.4285
27.4264 | -35.7019 | -38.4339
27.4325 | -35.7022 | -38.4392
27.4387 | -35.7026 | -38.4444
27.4448 | -35.7029 | -38.4497
27.4509 | -35.7033 | -38.4551
27.457 | -35.7036 | -38.4604
27.4631 | -35.704 | -38.4657
27.4692 | -35.7043 | -38.471
27.4753 | -35.7046 | -38.4761
27.4814 | -35.705 | -38.4815
27.4875 | -35.7053 | -38.4868
27.4936 | -35.7057 | -38.492
27.4997 | -35.706 | -38.4973
27.5058 | -35.7064 | -38.5025
27.5119 | -35.7067 | -38.5078
27.518 | -35.7071 | -38.513
27.5241 | -35.7074 | -38.5184
27.5302 | -35.7078 | -38.5238
27.5363 | -35.7081 | -38.5291
27.5424 | -35.7085 | -38.5343
27.5485 | -35.7088 | -38.5396
27.5546 | -35.7092 | -38.5448
27.5607 | -35.7095 | -38.5501
27.5668 | -35.7099 | -38.5555
27.5729 | -35.7102 | -38.5609
27.5791 | -35.7106 | -38.5661
27.5852 | -35.7109 | -38.5714
27.5913 | -35.7112 | -38.5766
27.5974 | -35.7116 | -38.5819
27.6035 | -35.7119 | -38.5873
27.6096 | -35.7122 | -38.5927
27.6157 | -35.7126 | -38.5978
27.6218 | -35.7129 | -38.6031
27.6279 | -35.7133 | -38.6085
27.634 | -35.7137 | -38.6138
27.6401 | -35.714 | -38.6192
27.6462 | -35.7143 | -38.6245
27.6523 | -35.7147 | -38.6298
27.6584 | -35.715 | -38.6352
27.6645 | -35.7154 | -38.6407
27.6706 | -35.7157 | -38.646
27.6767 | -35.716 | -38.6514
27.6828 | -35.7164 | -38.6565
27.6889 | -35.7168 | -38.6617
27.695 | -35.7171 | -38.667
27.7011 | -35.7175 | -38.6721
27.7072 | -35.7178 | -38.6773
27.7133 | -35.7182 | -38.6827
27.7194 | -35.7186 | -38.688
27.7256 | -35.7189 | -38.6934
27.7317 | -35.7193 | -38.6987
27.7378 | -35.7196 | -38.7039
27.7439 | -35.72 | -38.7094
27.75 | -35.7204 | -38.7146
27.7561 | -35.7207 | -38.7198
27.7622 | -35.721 | -38.7249
27.7683 | -35.7214 | -38.7301
27.7744 | -35.7217 | -38.7352
27.7805 | -35.7221 | -38.7407
27.7866 | -35.7224 | -38.746
27.7927 | -35.7228 | -38.7511
27.7988 | -35.7231 | -38.7563
27.8049 | -35.7235 | -38.7617
27.811 | -35.7238 | -38.7673
27.8171 | -35.7241 | -38.7726
27.8232 | -35.7245 | -38.7777
27.8293 | -35.7248 | -38.7827
27.8354 | -35.7252 | -38.7881
27.8415 | -35.7255 | -38.7933
27.8476 | -35.7258 | -38.7986
27.8537 | -35.7262 | -38.8038
27.8598 | -35.7265 | -38.8091
27.866 | -35.7268 | -38.8143
27.8721 | -35.7272 | -38.8196
27.8782 | -35.7275 | -38.8249
27.8843 | -35.7279 | -38.8302
27.8904 | -35.7282 | -38.8354
27.8965 | -35.7286 | -38.8408
27.9026 | -35.7289 | -38.846
27.9087 | -35.7293 | -38.8514
27.9148 | -35.7296 | -38.8567
27.9209 | -35.73 | -38.8619
27.927 | -35.7303 | -38.867
27.9331 | -35.7307 | -38.8723
27.9392 | -35.731 | -38.8775
27.9453 | -35.7314 | -38.8826
27.9514 | -35.7317 | -38.888
27.9575 | -35.7321 | -38.8933
27.9636 | -35.7324 | -38.8987
27.9697 | -35.7327 | -38.9038
27.9758 | -35.7331 | -38.9092
27.9819 | -35.7334 | -38.9145
27.988 | -35.7337 | -38.9198
27.9941 | -35.7341 | -38.9249
28.0002 | -35.7344 | -38.9303
28.0063 | -35.7347 | -38.9355
28.0125 | -35.7351 | -38.9405
28.0186 | -35.7354 | -38.9457
28.0247 | -35.7357 | -38.9508
28.0308 | -35.7361 | -38.956
28.0369 | -35.7364 | -38.9612
28.043 | -35.7367 | -38.9665
28.0491 | -35.7371 | -38.9716
28.0552 | -35.7374 | -38.9769
28.0613 | -35.7378 | -38.982
28.0674 | -35.7381 | -38.9874
28.0735 | -35.7385 | -38.9925

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4773 | -44.1623
21.4908 | -35.5159 | -35.9661
16.4948 | -34.8892 | -28.4259
12.6602 | -34.5192 | -22.0347
9.7171 | -34.2933 | -17.0369
7.4582 | -34.1385 | -12.9867
5.7244 | -34.0229 | -9.7084
4.3936 | -33.9571 | -7.0654
3.3722 | -33.9299 | -4.9846
2.5883 | -33.9261 | -3.4002
1.9866 | -33.9113 | -2.1392
1.5248 | -33.9175 | -1.1251
1.1703 | -33.9187 | -0.46211
1 | -33.9081 | -0.13158
0.89822 | -33.9014 | 0.14594
0.68941 | -33.8849 | 0.73742
0.52919 | -33.8773 | 1.192
0.40612 | -33.8783 | 1.8049
0.31171 | -33.8792 | 2.2622
0.23928 | -33.8664 | 2.9908
0.18358 | -33.8942 | 3.4064
0.14091 | -33.8739 | 4.2938
0.10819 | -33.9093 | 5.4431
0.083017 | -33.9178 | 6.9217
0.063705 | -33.9564 | 8.7932
0.048923 | -34.0177 | 11.2881
0.037479 | -34.1252 | 14.4676
0.028801 | -34.3032 | 18.4802
0.022078 | -34.611 | 23.8057
0.017023 | -35.0763 | 30.1982
0.013018 | -35.884 | 37.8834
0.010014 | -36.9982 | 45.7716

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003673 | -9.9486
0.007346 | -25.6732
0.011019 | -33.1785
0.014692 | -38.7402
0.018365 | -41.9076
0.022039 | -43.7084
0.025712 | -45.0615
0.029385 | -46.155
0.033058 | -47.2236
0.036731 | -47.3219
0.040404 | -48.7436
0.044077 | -51.5986
0.04775 | -52.7985
0.051423 | -53.7962
0.055096 | -54.6363
0.05877 | -55.7914
0.062443 | -55.9317
0.066116 | -56.9345
0.069789 | -58.364
0.073462 | -58.9585
0.077135 | -58.0584
0.080808 | -58.8956
0.084481 | -59.6482
0.088154 | -60.6665
0.091827 | -61.4787
0.0955 | -61.5362
0.099174 | -61.6462
0.10285 | -63.1731
0.10652 | -62.5571
0.11019 | -63.5614
0.11387 | -63.3049
0.11754 | -62.8429
0.12121 | -63.3826
0.12488 | -64.3304
0.12856 | -65.194
0.13223 | -65.1821
0.1359 | -66.018
0.13958 | -66.9809
0.14325 | -66.4534
0.14692 | -66.1541
0.1506 | -66.5804
0.15427 | -66.8735
0.15794 | -66.4961
0.16162 | -67.1185
0.16529 | -66.9362
0.16896 | -67.221
0.17264 | -67.6276
0.17631 | -67.6061
0.17998 | -68.8884
0.18366 | -67.7469
0.18733 | -68.2337
0.191 | -67.9649
0.19467 | -68.2411
0.19835 | -68.4378
0.20202 | -69.0253
0.20569 | -69.2092
0.20937 | -69.9293
0.21304 | -70.4105
0.21671 | -70.2089
0.22039 | -69.6328
0.22406 | -70.0321
0.22773 | -70.3726
0.2314 | -71.2704
0.23508 | -71.8684
0.23875 | -71.7807
0.24242 | -72.0649
0.2461 | -71.1185
0.24977 | -71.2956
0.25344 | -71.8363
0.25712 | -71.7063
0.26079 | -72.3874
0.26446 | -72.4848
0.26814 | -71.6225
0.27181 | -72.0237
0.27548 | -71.9927
0.27915 | -72.7632
0.28283 | -73.5687
0.2865 | -73.3711
0.29017 | -73.7488
0.29385 | -73.7079
0.29752 | -74.1894
0.30119 | -74.0335
0.30487 | -73.9893
0.30854 | -73.7635
0.31221 | -73.0436
0.31589 | -73.5098
0.31956 | -74.2929
0.32323 | -76.1301
0.32691 | -76.3904
0.33058 | -74.7288
0.33425 | -74.9341
0.33792 | -75.9299
0.3416 | -77.1103
0.34527 | -76.1274
0.34894 | -75.7793
0.35262 | -74.5082
0.35629 | -74.6732
0.35996 | -77.2055
0.36364 | -77.3796
0.36731 | -75.8909
0.37098 | -75.9217
0.37466 | -75.1477
0.37833 | -75.5932
0.382 | -76.3209
0.38567 | -76.1054
0.38935 | -76.3597
0.39302 | -75.5306
0.39669 | -76.2514
0.40037 | -76.4332
0.40404 | -76.9823
0.40771 | -78.1632
0.41139 | -77.7118
0.41506 | -76.2015
0.41873 | -76.4357
0.42241 | -77.7677
0.42608 | -78.2041
0.42975 | -78.165
0.43343 | -77.8744
0.4371 | -78.0555
0.44077 | -77.7675
0.44444 | -77.3468
0.44812 | -77.1538
0.45179 | -77.809
0.45546 | -77.9787
0.45914 | -77.637
0.46281 | -78.1568
0.46648 | -77.6875
0.47016 | -78.0998
0.47383 | -78.7003
0.4775 | -79.7902
0.48118 | -79.5484
0.48485 | -79.5979
0.48852 | -79.2636
0.49219 | -79.6882
0.49587 | -80.5345
0.49954 | -79.8897
0.50321 | -80.0187
0.50689 | -79.273
0.51056 | -78.5669
0.51423 | -78.4257
0.51791 | -79.7593
0.52158 | -80.3428
0.52525 | -80.2161
0.52893 | -80.5482
0.5326 | -80.5591
0.53627 | -80.4793
0.53995 | -80.4838
0.54362 | -80.829
0.54729 | -81.3017
0.55096 | -81.7034
0.55464 | -81.9484
0.55831 | -81.9436
0.56198 | -81.495
0.56566 | -80.8978
0.56933 | -80.5091
0.573 | -80.2227
0.57668 | -80.2798
0.58035 | -80.6707
0.58402 | -81.1101
0.58769 | -81.363
0.59137 | -81.368
0.59504 | -80.8334
0.59871 | -80.4355
0.60239 | -80.9609
0.60606 | -81.9378
0.60973 | -81.9617
0.61341 | -81.4201
0.61708 | -81.5543
0.62075 | -81.7648
0.62443 | -81.7432
0.6281 | -82.0541
0.63177 | -82.512
0.63545 | -82.6876
0.63912 | -82.2984
0.64279 | -81.6809
0.64646 | -81.5584
0.65014 | -82.1323
0.65381 | -82.5482
0.65748 | -82.5877
0.66116 | -82.9176
0.66483 | -83.2127
0.6685 | -83.0088
0.67218 | -82.6105
0.67585 | -82.1888
0.67952 | -82.2887
0.6832 | -83.0752
0.68687 | -83.3594
0.69054 | -82.8816
0.69422 | -82.6485
0.69789 | -82.5296
0.70156 | -82.2954
0.70523 | -82.6339
0.70891 | -83.4483
0.71258 | -83.2979
0.71625 | -83.0305
0.71993 | -83.3438
0.7236 | -83.7379
0.72727 | -83.9398
0.73095 | -83.8905
0.73462 | -83.659
0.73829 | -83.4419
0.74196 | -83.6308
0.74564 | -84.1147
0.74931 | -84.4353
0.75298 | -84.6043
0.75666 | -84.3709
0.76033 | -84.3271
0.764 | -84.1082
0.76768 | -83.7223
0.77135 | -83.7834
0.77502 | -83.9649
0.7787 | -84.1024
0.78237 | -84.065
0.78604 | -84.0952
0.78971 | -84.1567
0.79339 | -83.786
0.79706 | -83.6136
0.80073 | -83.58
0.80441 | -83.8817
0.80808 | -84.6266
0.81175 | -85.5033
0.81543 | -85.765
0.8191 | -85.5674
0.82277 | -85.6447
0.82645 | -85.4848
0.83012 | -85.1332
0.83379 | -85.1156
0.83747 | -85.5613
0.84114 | -85.7345
0.84481 | -85.0952
0.84849 | -84.4879
0.85216 | -84.2446
0.85583 | -84.7624
0.8595 | -85.8636
0.86318 | -85.9892
0.86685 | -84.924
0.87052 | -84.4024
0.8742 | -84.6384
0.87787 | -84.9353
0.88154 | -85.0972
0.88522 | -84.8288
0.88889 | -84.5505
0.89256 | -84.9194
0.89624 | -85.6146
0.89991 | -85.8734
0.90358 | -85.6407
0.90725 | -84.8471
0.91093 | -84.165
0.9146 | -84.2527
0.91827 | -84.8531
0.92195 | -85.4675
0.92562 | -85.9826
0.92929 | -86.4068
0.93297 | -86.32
0.93664 | -86.3457
0.94031 | -86.569
0.94398 | -86.5716
0.94766 | -86.3085
0.95133 | -86.2306
0.955 | -86.6687
0.95868 | -87.076
0.96235 | -87.1036
0.96602 | -86.9191
0.9697 | -86.8658
0.97337 | -87.0082
0.97704 | -86.9689
0.98072 | -86.6822
0.98439 | -86.7096
0.98806 | -87.0015
0.99174 | -86.9672
0.99541 | -86.5712
0.99908 | -86.039
1.0028 | -85.8404
1.0064 | -85.8768
1.0101 | -86.014
1.0138 | -86.3979
1.0174 | -86.7762
1.0211 | -86.7129
1.0248 | -86.6845
1.0285 | -86.7792
1.0321 | -86.8251
1.0358 | -86.837
1.0395 | -87.1064
1.0432 | -87.5449
1.0468 | -87.6567
1.0505 | -87.9802
1.0542 | -88.6578
1.0579 | -87.9788
1.0615 | -87.4311
1.0652 | -87.1213
1.0689 | -87.187
1.0725 | -87.5822
1.0762 | -88.0237
1.0799 | -88.0371
1.0836 | -87.639
1.0872 | -87.2827
1.0909 | -87.211
1.0946 | -87.2768
1.0983 | -87.3117
1.1019 | -87.3796
1.1056 | -87.4789
1.1093 | -87.4332
1.1129 | -87.1849
1.1166 | -86.8854
1.1203 | -86.7723
1.124 | -86.9801
1.1276 | -87.3441
1.1313 | -87.571
1.135 | -87.6396
1.1387 | -87.6121
1.1423 | -87.5128
1.146 | -87.4653
1.1497 | -87.486
1.1534 | -87.4926
1.157 | -87.507
1.1607 | -87.7005
1.1644 | -88.1013
1.168 | -88.353
1.1717 | -88.2232
1.1754 | -87.9534
1.1791 | -87.7839
1.1827 | -87.7858
1.1864 | -87.8951
1.1901 | -88.0838
1.1938 | -88.2522
1.1974 | -88.2749
1.2011 | -88.3467
1.2048 | -88.619
1.2084 | -88.8739
1.2121 | -88.8704
1.2158 | -88.6044
1.2195 | -88.3091
1.2231 | -88.1649
1.2268 | -88.1659
1.2305 | -88.2077
1.2342 | -88.2132
1.2378 | -88.1572
1.2415 | -88.0937
1.2452 | -88.147
1.2489 | -88.2522
1.2525 | -88.2087
1.2562 | -88.0354
1.2599 | -87.8377
1.2635 | -87.6367
1.2672 | -87.5291
1.2709 | -87.7479
1.2746 | -88.3946
1.2782 | -89.1927
1.2819 | -89.5898
1.2856 | -89.3804
1.2893 | -89.0098
1.2929 | -88.8643
1.2966 | -88.9672
1.3003 | -89.1112
1.3039 | -89.0968
1.3076 | -88.9768
1.3113 | -88.9474
1.315 | -89.1707
1.3186 | -89.6079
1.3223 | -89.9193
1.326 | -89.8611
1.3297 | -89.5289
1.3333 | -89.0526
1.337 | -88.6248
1.3407 | -88.4737
1.3444 | -88.6396
1.348 | -88.9402
1.3517 | -89.1158
1.3554 | -89.0212
1.359 | -88.8581
1.3627 | -88.8991
1.3664 | -89.029
1.3701 | -89.009
1.3737 | -89.1039
1.3774 | -89.5864
1.3811 | -90.045
1.3848 | -89.9257
1.3884 | -89.4814
1.3921 | -89.1348
1.3958 | -88.976
1.3994 | -88.883
1.4031 | -88.8867
1.4068 | -89.2404
1.4105 | -89.8789
1.4141 | -90.3038
1.4178 | -90.4001
1.4215 | -90.4656
1.4252 | -90.4896
1.4288 | -90.436
1.4325 | -90.4723
1.4362 | -90.5692
1.4399 | -90.6253
1.4435 | -90.6509
1.4472 | -90.5048
1.4509 | -90.1431
1.4545 | -89.8829
1.4582 | -89.9079
1.4619 | -90.0363
1.4656 | -90.1264
1.4692 | -90.3309
1.4729 | -90.7252
1.4766 | -91.0084
1.4803 | -90.924
1.4839 | -90.7018
1.4876 | -90.6592
1.4913 | -90.7723
1.4949 | -90.7699
1.4986 | -90.5043
1.5023 | -90.0628
1.506 | -89.5952
1.5096 | -89.2481
1.5133 | -89.1039
1.517 | -89.1883
1.5207 | -89.555
1.5243 | -90.0913
1.528 | -90.4579
1.5317 | -90.5361
1.5354 | -90.4934
1.539 | -90.4486
1.5427 | -90.4654
1.5464 | -90.4529
1.55 | -90.3534
1.5537 | -90.3288
1.5574 | -90.4505
1.5611 | -90.5857
1.5647 | -90.6975
1.5684 | -90.9266
1.5721 | -91.2998
1.5758 | -91.5872
1.5794 | -91.667
1.5831 | -91.7
1.5868 | -91.6813
1.5904 | -91.6282
1.5941 | -91.5605
1.5978 | -91.535
1.6015 | -91.5348
1.6051 | -91.4873
1.6088 | -91.3563
1.6125 | -91.1461
1.6162 | -90.949
1.6198 | -90.9091
1.6235 | -91.0375
1.6272 | -91.1838
1.6309 | -91.1606
1.6345 | -91.0116
1.6382 | -90.9014
1.6419 | -90.9109
1.6455 | -91.0162
1.6492 | -91.1604
1.6529 | -91.2687
1.6566 | -91.2478
1.6602 | -91.0789
1.6639 | -90.8996
1.6676 | -90.8621
1.6713 | -90.9924
1.6749 | -91.2067
1.6786 | -91.4037
1.6823 | -91.5124
1.686 | -91.5326
1.6896 | -91.5056
1.6933 | -91.5171
1.697 | -91.6078
1.7006 | -91.7316
1.7043 | -91.874
1.708 | -92.0619
1.7117 | -92.2423
1.7153 | -92.2856
1.719 | -92.1514
1.7227 | -91.9943
1.7264 | -91.9231
1.73 | -91.9571
1.7337 | -92.0755
1.7374 | -92.1803
1.741 | -92.1622
1.7447 | -92.0365
1.7484 | -91.9134
1.7521 | -91.8976
1.7557 | -91.9978
1.7594 | -92.1802
1.7631 | -92.394
1.7668 | -92.5553
1.7704 | -92.5682
1.7741 | -92.3788
1.7778 | -92.0504
1.7815 | -91.7089
1.7851 | -91.4306
1.7888 | -91.2007
1.7925 | -91.0199
1.7961 | -90.9871
1.7998 | -91.193
1.8035 | -91.5672
1.8072 | -91.8936
1.8108 | -92.0098
1.8145 | -91.9111
1.8182 | -91.7006
1.8219 | -91.5339
1.8255 | -91.4629
1.8292 | -91.423
1.8329 | -91.3977
1.8365 | -91.4672
1.8402 | -91.6903
1.8439 | -92.0078
1.8476 | -92.3326
1.8512 | -92.5931
1.8549 | -92.7252
1.8586 | -92.684
1.8623 | -92.4908
1.8659 | -92.2558
1.8696 | -92.1543
1.8733 | -92.2532
1.877 | -92.4958
1.8806 | -92.7227
1.8843 | -92.8233
1.888 | -92.7701
1.8916 | -92.5936
1.8953 | -92.3627
1.899 | -92.1434
1.9027 | -91.9836
1.9063 | -91.9173
1.91 | -91.9003
1.9137 | -91.8485
1.9174 | -91.7812
1.921 | -91.8442
1.9247 | -92.1016
1.9284 | -92.4621
1.932 | -92.7777
1.9357 | -92.9512
1.9394 | -92.9499
1.9431 | -92.7853
1.9467 | -92.5414
1.9504 | -92.3177
1.9541 | -92.1878
1.9578 | -92.1892
1.9614 | -92.2962
1.9651 | -92.4376
1.9688 | -92.5413
1.9725 | -92.5768
1.9761 | -92.5748
1.9798 | -92.5736
1.9835 | -92.5485
1.9871 | -92.4858
1.9908 | -92.4233
1.9945 | -92.4139
1.9982 | -92.4789
2.0018 | -92.5939
2.0055 | -92.7587
2.0092 | -92.9601
2.0129 | -93.1571
2.0165 | -93.2624
2.0202 | -93.2521
2.0239 | -93.2029
2.0275 | -93.1877
2.0312 | -93.1956
2.0349 | -93.2035
2.0386 | -93.238
2.0422 | -93.3324
2.0459 | -93.4328
2.0496 | -93.4318
2.0533 | -93.2992
2.0569 | -93.1595
2.0606 | -93.1493
2.0643 | -93.2613
2.068 | -93.3707
2.0716 | -93.3626
2.0753 | -93.2657
2.079 | -93.2104
2.0826 | -93.29
2.0863 | -93.4535
2.09 | -93.5582
2.0937 | -93.5173
2.0973 | -93.3926
2.101 | -93.2426
2.1047 | -93.0561
2.1084 | -92.8207
2.112 | -92.6165
2.1157 | -92.5111
2.1194 | -92.5591
2.123 | -92.5974
2.1267 | -92.6773
2.1304 | -92.7937
2.1341 | -92.9542
2.1377 | -93.1452
2.1414 | -93.341
2.1451 | -93.5064
2.1488 | -93.6106
2.1524 | -93.6786
2.1561 | -93.7259
2.1598 | -93.7579
2.1635 | -93.7688
2.1671 | -93.7475
2.1708 | -93.6773
2.1745 | -93.5521
2.1781 | -93.3938
2.1818 | -93.25
2.1855 | -93.2017
2.1892 | -93.2948
2.1928 | -93.5104
2.1965 | -93.7807
2.2002 | -94.013
2.2039 | -94.1403
2.2075 | -94.1397
2.2112 | -94.0684
2.2149 | -93.9957
2.2185 | -93.9533
2.2222 | -93.9248
2.2259 | -93.8412
2.2296 | -93.6931
2.2332 | -93.5164
2.2369 | -93.3635
2.2406 | -93.2787
2.2443 | -93.2592
2.2479 | -93.2707
2.2516 | -93.27
2.2553 | -93.2665
2.259 | -93.3261
2.2626 | -93.4835
2.2663 | -93.7433
2.27 | -94.0344
2.2736 | -94.2501
2.2773 | -94.3355
2.281 | -94.3132
2.2847 | -94.2881
2.2883 | -94.3291
2.292 | -94.4513
2.2957 | -94.5957
2.2994 | -94.6733
2.303 | -94.6462
2.3067 | -94.4948
2.3104 | -94.2535
2.314 | -94.0058
2.3177 | -93.7954
2.3214 | -93.6499
2.3251 | -93.5442
2.3287 | -93.4811
2.3324 | -93.4799
2.3361 | -93.5508
2.3398 | -93.6696
2.3434 | -93.8033
2.3471 | -93.9168
2.3508 | -93.9838
2.3545 | -93.9709
2.3581 | -93.8834
2.3618 | -93.7712
2.3655 | -93.7002
2.3691 | -93.7054
2.3728 | -93.7879
2.3765 | -93.9363
2.3802 | -94.1093
2.3838 | -94.2765
2.3875 | -94.4312
2.3912 | -94.5691
2.3949 | -94.7038
2.3985 | -94.8265
2.4022 | -94.9324
2.4059 | -95.0038
2.4095 | -95.0164
2.4132 | -94.943
2.4169 | -94.7811
2.4206 | -94.57
2.4242 | -94.3616
2.4279 | -94.1991
2.4316 | -94.0967
2.4353 | -94.0393
2.4389 | -94.0023
2.4426 | -93.9686
2.4463 | -93.9523
2.45 | -93.9618
2.4536 | -93.9865
2.4573 | -94.0199
2.461 | -94.0673
2.4646 | -94.1397
2.4683 | -94.254
2.472 | -94.4076
2.4757 | -94.583
2.4793 | -94.7431
2.483 | -94.816
2.4867 | -94.7758
2.4904 | -94.6819
2.494 | -94.6289
2.4977 | -94.6573
2.5014 | -94.7494
2.5051 | -94.8527
2.5087 | -94.899
2.5124 | -94.8361
2.5161 | -94.6916
2.5197 | -94.5428
2.5234 | -94.4546
2.5271 | -94.4583
2.5308 | -94.518
2.5344 | -94.5878
2.5381 | -94.6192
2.5418 | -94.568
2.5455 | -94.4583
2.5491 | -94.3809
2.5528 | -94.4121
2.5565 | -94.5413
2.5601 | -94.7056
2.5638 | -94.8342
2.5675 | -94.8958
2.5712 | -94.8882
2.5748 | -94.7913
2.5785 | -94.5963
2.5822 | -94.3422
2.5859 | -94.0914
2.5895 | -93.9377
2.5932 | -93.9621
2.5969 | -94.2056
2.6006 | -94.6412
2.6042 | -95.1551
2.6079 | -95.5955
2.6116 | -95.8536
2.6152 | -95.9067
2.6189 | -95.8052
2.6226 | -95.6271
2.6263 | -95.4201
2.6299 | -95.2157
2.6336 | -95.046
2.6373 | -94.9267
2.641 | -94.8638
2.6446 | -94.8508
2.6483 | -94.8361
2.652 | -94.81
2.6556 | -94.7853
2.6593 | -94.7683
2.663 | -94.7639
2.6667 | -94.7727
2.6703 | -94.7807
2.674 | -94.7845
2.6777 | -94.7825
2.6814 | -94.7798
2.685 | -94.7864
2.6887 | -94.8197
2.6924 | -94.8854
2.6961 | -94.9825
2.6997 | -95.1006
2.7034 | -95.22
2.7071 | -95.3399
2.7107 | -95.4544
2.7144 | -95.561
2.7181 | -95.6616
2.7218 | -95.7404
2.7254 | -95.8135
2.7291 | -95.8876
2.7328 | -95.9601
2.7365 | -96.0106
2.7401 | -96.0181
2.7438 | -95.9795
2.7475 | -95.8986
2.7511 | -95.7852
2.7548 | -95.6475
2.7585 | -95.5185
2.7622 | -95.4203
2.7658 | -95.363
2.7695 | -95.3541
2.7732 | -95.366
2.7769 | -95.3568
2.7805 | -95.3185
2.7842 | -95.274
2.7879 | -95.2623
2.7916 | -95.2975
2.7952 | -95.3728
2.7989 | -95.4452
2.8026 | -95.4821
2.8062 | -95.4557
2.8099 | -95.3831
2.8136 | -95.3116
2.8173 | -95.2643
2.8209 | -95.2603
2.8246 | -95.2995
2.8283 | -95.3427
2.832 | -95.3663
2.8356 | -95.3524
2.8393 | -95.3169
2.843 | -95.2971
2.8466 | -95.3347
2.8503 | -95.4227
2.854 | -95.5412
2.8577 | -95.6585
2.8613 | -95.7417
2.865 | -95.7689
2.8687 | -95.7496
2.8724 | -95.7122
2.876 | -95.6651
2.8797 | -95.624
2.8834 | -95.6194
2.8871 | -95.652
2.8907 | -95.7173
2.8944 | -95.8078
2.8981 | -95.9121
2.9017 | -96.0103
2.9054 | -96.0903
2.9091 | -96.1563
2.9128 | -96.2098
2.9164 | -96.2439
2.9201 | -96.2244
2.9238 | -96.1621
2.9275 | -96.0626
2.9311 | -95.9509
2.9348 | -95.8764
2.9385 | -95.8627
2.9421 | -95.8994
2.9458 | -95.9657
2.9495 | -96.0232
2.9532 | -96.0318
2.9568 | -95.9925
2.9605 | -95.9371
2.9642 | -95.8961
2.9679 | -95.9096
2.9715 | -95.9794
2.9752 | -96.0831
2.9789 | -96.1988
2.9826 | -96.2993
2.9862 | -96.3659
2.9899 | -96.4034
2.9936 | -96.4149
2.9972 | -96.4088
3.0009 | -96.3964
3.0046 | -96.3875
3.0083 | -96.3959
3.0119 | -96.4195
3.0156 | -96.4635
3.0193 | -96.5239
3.023 | -96.6059
3.0266 | -96.7102
3.0303 | -96.8569
3.034 | -97.0447
3.0376 | -97.2626
3.0413 | -97.4637
3.045 | -97.5994
3.0487 | -97.6366
3.0523 | -97.5756
3.056 | -97.4202
3.0597 | -97.2202
3.0634 | -97.0048
3.067 | -96.7923
3.0707 | -96.5885
3.0744 | -96.4043
3.0781 | -96.2542
3.0817 | -96.1597
3.0854 | -96.1249
3.0891 | -96.1574
3.0927 | -96.2309
3.0964 | -96.3266
3.1001 | -96.4361
3.1038 | -96.5421
3.1074 | -96.6332
3.1111 | -96.6884
3.1148 | -96.7035
3.1185 | -96.6739
3.1221 | -96.6184
3.1258 | -96.577
3.1295 | -96.5763
3.1332 | -96.6102
3.1368 | -96.6653
3.1405 | -96.7263
3.1442 | -96.7625
3.1478 | -96.7634
3.1515 | -96.7419
3.1552 | -96.7077
3.1589 | -96.6921
3.1625 | -96.7075
3.1662 | -96.7547
3.1699 | -96.8134
3.1736 | -96.8724
3.1772 | -96.9306
3.1809 | -96.9865
3.1846 | -97.0378
3.1882 | -97.0719
3.1919 | -97.0796
3.1956 | -97.0644
3.1993 | -97.0323
3.2029 | -96.9741
3.2066 | -96.8992
3.2103 | -96.8176
3.214 | -96.738
3.2176 | -96.6782
3.2213 | -96.6458
3.225 | -96.646
3.2287 | -96.6774
3.2323 | -96.7363
3.236 | -96.8182
3.2397 | -96.9222
3.2433 | -97.0234
3.247 | -97.1104
3.2507 | -97.1763
3.2544 | -97.2339
3.258 | -97.2958
3.2617 | -97.3552
3.2654 | -97.4281
3.2691 | -97.5074
3.2727 | -97.5723
3.2764 | -97.6163
3.2801 | -97.6244
3.2837 | -97.6161
3.2874 | -97.5961
3.2911 | -97.5665
3.2948 | -97.535
3.2984 | -97.5137
3.3021 | -97.4796
3.3058 | -97.4324
3.3095 | -97.3662
3.3131 | -97.3007
3.3168 | -97.2394
3.3205 | -97.1991
3.3242 | -97.188
3.3278 | -97.1874
3.3315 | -97.1837
3.3352 | -97.1719
3.3388 | -97.1517
3.3425 | -97.1232
3.3462 | -97.0997
3.3499 | -97.0786
3.3535 | -97.0688
3.3572 | -97.0893
3.3609 | -97.1408
3.3646 | -97.2212
3.3682 | -97.3158
3.3719 | -97.4143
3.3756 | -97.4918
3.3792 | -97.5302
3.3829 | -97.5408
3.3866 | -97.5151
3.3903 | -97.4505
3.3939 | -97.35
3.3976 | -97.2477
3.4013 | -97.1562
3.405 | -97.1002
3.4086 | -97.0913
3.4123 | -97.1311
3.416 | -97.2145
3.4197 | -97.3268
3.4233 | -97.4465
3.427 | -97.5512
3.4307 | -97.6292
3.4343 | -97.6731
3.438 | -97.6631
3.4417 | -97.6238
3.4454 | -97.5706
3.449 | -97.5184
3.4527 | -97.4663
3.4564 | -97.4132
3.4601 | -97.355
3.4637 | -97.2921
3.4674 | -97.2316
3.4711 | -97.1853
3.4747 | -97.1582
3.4784 | -97.1488
3.4821 | -97.1504
3.4858 | -97.1474
3.4894 | -97.1495
3.4931 | -97.1683
3.4968 | -97.2061
3.5005 | -97.2892
3.5041 | -97.4224
3.5078 | -97.5867
3.5115 | -97.7455
3.5152 | -97.8583
3.5188 | -97.9126
3.5225 | -97.9132
3.5262 | -97.8771
3.5298 | -97.825
3.5335 | -97.7599
3.5372 | -97.6976
3.5409 | -97.6588
3.5445 | -97.6319
3.5482 | -97.6077
3.5519 | -97.5756
3.5556 | -97.5183
3.5592 | -97.4282
3.5629 | -97.3175
3.5666 | -97.1953
3.5702 | -97.0881
3.5739 | -97.0191
3.5776 | -96.9867
3.5813 | -96.9889
3.5849 | -97.0207
3.5886 | -97.0755
3.5923 | -97.1388
3.596 | -97.2094
3.5996 | -97.312
3.6033 | -97.4449
3.607 | -97.5949
3.6107 | -97.7463
3.6143 | -97.8818
3.618 | -97.9856
3.6217 | -98.0492
3.6253 | -98.0363
3.629 | -97.9556
3.6327 | -97.8425
3.6364 | -97.7153
3.64 | -97.5932
3.6437 | -97.5031
3.6474 | -97.4685
3.6511 | -97.499
3.6547 | -97.5823
3.6584 | -97.7014
3.6621 | -97.8331
3.6657 | -97.9488
3.6694 | -98.0323
3.6731 | -98.0876
3.6768 | -98.1139
3.6804 | -98.1415
3.6841 | -98.1629
3.6878 | -98.1902
3.6915 | -98.2268
3.6951 | -98.2754
3.6988 | -98.3275
3.7025 | -98.3659
3.7062 | -98.3306
3.7098 | -98.3028
3.7135 | -98.2535
3.7172 | -98.1903
3.7208 | -98.1323
3.7245 | -98.0829
3.7282 | -98.0612
3.7319 | -98.0716
3.7355 | -98.1233
3.7392 | -98.2037
3.7429 | -98.3152
3.7466 | -98.4279
3.7502 | -98.527
3.7539 | -98.5889
3.7576 | -98.6069
3.7612 | -98.5946
3.7649 | -98.5552
3.7686 | -98.5005
3.7723 | -98.4374
3.7759 | -98.3716
3.7796 | -98.309
3.7833 | -98.243
3.787 | -98.199
3.7906 | -98.1708
3.7943 | -98.1587
3.798 | -98.1636
3.8017 | -98.1804
3.8053 | -98.2098
3.809 | -98.2593
3.8127 | -98.3225
3.8163 | -98.3919
3.82 | -98.4446
3.8237 | -98.4663
3.8274 | -98.4618
3.831 | -98.4296
3.8347 | -98.3789
3.8384 | -98.3277
3.8421 | -98.2964
3.8457 | -98.2849
3.8494 | -98.2919
3.8531 | -98.3236
3.8567 | -98.3761
3.8604 | -98.4619
3.8641 | -98.5812
3.8678 | -98.7199
3.8714 | -98.8719
3.8751 | -99.0107
3.8788 | -99.1099
3.8825 | -99.1579
3.8861 | -99.1592
3.8898 | -99.1174
3.8935 | -99.063
3.8972 | -99.0146
3.9008 | -98.9732
3.9045 | -98.9448
3.9082 | -98.9332
3.9118 | -98.9286
3.9155 | -98.9285
3.9192 | -98.9316
3.9229 | -98.9287
3.9265 | -98.9429
3.9302 | -98.9745
3.9339 | -99.0092
3.9376 | -99.0367
3.9412 | -99.0515
3.9449 | -99.0414
3.9486 | -99.0037
3.9522 | -98.9559
3.9559 | -98.9006
3.9596 | -98.855
3.9633 | -98.8369
3.9669 | -98.8353
3.9706 | -98.8455
3.9743 | -98.8568
3.978 | -98.8662
3.9816 | -98.8609
3.9853 | -98.8391
3.989 | -98.8
3.9927 | -98.7651
3.9963 | -98.7366
4 | -98.7275
4.0037 | -98.7344
4.0073 | -98.7587
4.011 | -98.7989
4.0147 | -98.8443
4.0184 | -98.8907
4.022 | -98.926
4.0257 | -98.9458
4.0294 | -98.9571
4.0331 | -98.9687
4.0367 | -98.9934
4.0404 | -99.0271
4.0441 | -99.0734
4.0477 | -99.1218
4.0514 | -99.1664
4.0551 | -99.1985
4.0588 | -99.2332
4.0624 | -99.2697
4.0661 | -99.3219
4.0698 | -99.3892
4.0735 | -99.469
4.0771 | -99.5562
4.0808 | -99.6318
4.0845 | -99.6769
4.0882 | -99.6749
4.0918 | -99.6263
4.0955 | -99.5505
4.0992 | -99.4543
4.1028 | -99.3641
4.1065 | -99.2798
4.1102 | -99.2115
4.1139 | -99.1581
4.1175 | -99.1111
4.1212 | -99.0487
4.1249 | -98.9794
4.1286 | -98.9001
4.1322 | -98.8202
4.1359 | -98.743
4.1396 | -98.6734
4.1433 | -98.6173
4.1469 | -98.5887
4.1506 | -98.5996
4.1543 | -98.6416
4.1579 | -98.7098
4.1616 | -98.7938
4.1653 | -98.8809
4.169 | -98.9662
4.1726 | -99.042
4.1763 | -99.1111
4.18 | -99.1813
4.1837 | -99.2549
4.1873 | -99.3065
4.191 | -99.3425
4.1947 | -99.358
4.1983 | -99.3481
4.202 | -99.3402
4.2057 | -99.3376
4.2094 | -99.3355
4.213 | -99.3339
4.2167 | -99.3395
4.2204 | -99.3377
4.2241 | -99.3375
4.2277 | -99.337
4.2314 | -99.3504
4.2351 | -99.3623
4.2388 | -99.3699
4.2424 | -99.3712
4.2461 | -99.3655
4.2498 | -99.3539
4.2534 | -99.3405
4.2571 | -99.3117
4.2608 | -99.2732
4.2645 | -99.2191
4.2681 | -99.1435
4.2718 | -99.0634
4.2755 | -98.9884
4.2792 | -98.9179
4.2828 | -98.8656
4.2865 | -98.8428
4.2902 | -98.8408
4.2938 | -98.8552
4.2975 | -98.881
4.3012 | -98.9181
4.3049 | -98.9648
4.3085 | -99.0017
4.3122 | -99.0354
4.3159 | -99.0555
4.3196 | -99.0558
4.3232 | -99.0368
4.3269 | -99.0116
4.3306 | -98.9944
4.3343 | -99.0049
4.3379 | -99.0383
4.3416 | -99.0997
4.3453 | -99.1903
4.3489 | -99.3043
4.3526 | -99.4307
4.3563 | -99.5573
4.36 | -99.6754
4.3636 | -99.7744
4.3673 | -99.8447
4.371 | -99.8932
4.3747 | -99.9254
4.3783 | -99.9499
4.382 | -99.9564
4.3857 | -99.9571
4.3893 | -99.9606
4.393 | -99.9664
4.3967 | -99.9592
4.4004 | -99.9295
4.404 | -99.8853
4.4077 | -99.8426
4.4114 | -99.8071
4.4151 | -99.7728
4.4187 | -99.7588
4.4224 | -99.7712
4.4261 | -99.8066
4.4298 | -99.8642
4.4334 | -99.9275
4.4371 | -99.9889
4.4408 | -100.0402
4.4444 | -100.0634
4.4481 | -100.063
4.4518 | -100.0462
4.4555 | -100.0031
4.4591 | -99.9362
4.4628 | -99.858
4.4665 | -99.7906
4.4702 | -99.7372
4.4738 | -99.696
4.4775 | -99.6721
4.4812 | -99.6597
4.4848 | -99.6719
4.4885 | -99.691
4.4922 | -99.7155
4.4959 | -99.752
4.4995 | -99.7958
4.5032 | -99.851
4.5069 | -99.9064
4.5106 | -99.9569
4.5142 | -100.0028
4.5179 | -100.0384
4.5216 | -100.0671
4.5253 | -100.0995
4.5289 | -100.1359
4.5326 | -100.1711
4.5363 | -100.2055
4.5399 | -100.2501
4.5436 | -100.3093
4.5473 | -100.3673
4.551 | -100.4162
4.5546 | -100.4471
4.5583 | -100.4562
4.562 | -100.4368
4.5657 | -100.39
4.5693 | -100.3195
4.573 | -100.2323
4.5767 | -100.1326
4.5803 | -100.0374
4.584 | -99.9533
4.5877 | -99.8795
4.5914 | -99.8121
4.595 | -99.7563
4.5987 | -99.7123
4.6024 | -99.6825
4.6061 | -99.6611
4.6097 | -99.635
4.6134 | -99.6091
4.6171 | -99.5973
4.6208 | -99.5949
4.6244 | -99.5921
4.6281 | -99.5892
4.6318 | -99.589
4.6354 | -99.5948
4.6391 | -99.6058
4.6428 | -99.6223
4.6465 | -99.6422
4.6501 | -99.657
4.6538 | -99.6726
4.6575 | -99.6845
4.6612 | -99.686
4.6648 | -99.6804
4.6685 | -99.6764
4.6722 | -99.6768
4.6758 | -99.6882
4.6795 | -99.7196
4.6832 | -99.7636
4.6869 | -99.81
4.6905 | -99.8733
4.6942 | -99.9506
4.6979 | -100.0357
4.7016 | -100.1204
4.7052 | -100.2022
4.7089 | -100.2746
4.7126 | -100.3245
4.7163 | -100.3519
4.7199 | -100.356
4.7236 | -100.3362
4.7273 | -100.3123
4.7309 | -100.2902
4.7346 | -100.2695
4.7383 | -100.2518
4.742 | -100.2462
4.7456 | -100.2551
4.7493 | -100.2732
4.753 | -100.3023
4.7567 | -100.331
4.7603 | -100.3377
4.764 | -100.3295
4.7677 | -100.2736
4.7713 | -100.2362
4.775 | -100.1889
4.7787 | -100.1439
4.7824 | -100.1017
4.786 | -100.0747
4.7897 | -100.0738
4.7934 | -100.0825
4.7971 | -100.1034
4.8007 | -100.1319
4.8044 | -100.1617
4.8081 | -100.1899
4.8118 | -100.2098
4.8154 | -100.2139
4.8191 | -100.1962
4.8228 | -100.1653
4.8264 | -100.1169
4.8301 | -100.0599
4.8338 | -100.01
4.8375 | -99.9556
4.8411 | -99.9174
4.8448 | -99.8973
4.8485 | -99.9052
4.8522 | -99.937
4.8558 | -99.9883
4.8595 | -100.0662
4.8632 | -100.1582
4.8669 | -100.2547
4.8705 | -100.3653
4.8742 | -100.4778
4.8779 | -100.5762
4.8815 | -100.6643
4.8852 | -100.7356
4.8889 | -100.781
4.8926 | -100.8133
4.8962 | -100.8269
4.8999 | -100.816
4.9036 | -100.7883
4.9073 | -100.7477
4.9109 | -100.6943
4.9146 | -100.6476
4.9183 | -100.6097
4.9219 | -100.5825
4.9256 | -100.5681
4.9293 | -100.5661
4.933 | -100.5689
4.9366 | -100.5721
4.9403 | -100.5763
4.944 | -100.5739
4.9477 | -100.5618
4.9513 | -100.5475
4.955 | -100.5471
4.9587 | -100.5528
4.9624 | -100.5589
4.966 | -100.5801
4.9697 | -100.5974
4.9734 | -100.6256
4.977 | -100.6562
4.9807 | -100.6821
4.9844 | -100.7019
4.9881 | -100.7172
4.9917 | -100.7226
4.9954 | -100.7123
4.9991 | -100.6982
5.0028 | -100.6779
5.0064 | -100.658
5.0101 | -100.6475
5.0138 | -100.6404
5.0174 | -100.6396
5.0211 | -100.6375
5.0248 | -100.6333
5.0285 | -100.624
5.0321 | -100.5973
5.0358 | -100.5622
5.0395 | -100.5314
5.0432 | -100.4958
5.0468 | -100.4639
5.0505 | -100.435
5.0542 | -100.42
5.0579 | -100.4178
5.0615 | -100.4249
5.0652 | -100.4399
5.0689 | -100.4694
5.0725 | -100.5122
5.0762 | -100.5669
5.0799 | -100.6308
5.0836 | -100.7101
5.0872 | -100.7953
5.0909 | -100.8887
5.0946 | -100.9841
5.0983 | -101.0706
5.1019 | -101.1379
5.1056 | -101.1808
5.1093 | -101.1972
5.1129 | -101.185
5.1166 | -101.1446
5.1203 | -101.0959
5.124 | -101.0359
5.1276 | -100.9781
5.1313 | -100.9327
5.135 | -100.8972
5.1387 | -100.8716
5.1423 | -100.8582
5.146 | -100.8577
5.1497 | -100.8715
5.1534 | -100.8932
5.157 | -100.9187
5.1607 | -100.931
5.1644 | -100.9267
5.168 | -100.914
5.1717 | -100.8805
5.1754 | -100.8367
5.1791 | -100.7878
5.1827 | -100.7395
5.1864 | -100.7016
5.1901 | -100.6808
5.1938 | -100.6698
5.1974 | -100.6625
5.2011 | -100.6759
5.2048 | -100.6953
5.2084 | -100.7051
5.2121 | -100.7163
5.2158 | -100.7245
5.2195 | -100.7237
5.2231 | -100.7149
5.2268 | -100.7049
5.2305 | -100.6949
5.2342 | -100.6894
5.2378 | -100.698
5.2415 | -100.719
5.2452 | -100.7382
5.2489 | -100.7675
5.2525 | -100.8031
5.2562 | -100.8356
5.2599 | -100.867
5.2635 | -100.8902
5.2672 | -100.8998
5.2709 | -100.8985
5.2746 | -100.8877
5.2782 | -100.8555
5.2819 | -100.8193
5.2856 | -100.795
5.2893 | -100.7736
5.2929 | -100.7588
5.2966 | -100.7793
5.3003 | -100.7933
5.3039 | -100.8177
5.3076 | -100.8512
5.3113 | -100.8872
5.315 | -100.9294
5.3186 | -100.9707
5.3223 | -101.0161
5.326 | -101.0552
5.3297 | -101.0824
5.3333 | -101.1021
5.337 | -101.1119
5.3407 | -101.1134
5.3444 | -101.1106
5.348 | -101.1064
5.3517 | -101.107
5.3554 | -101.1019
5.359 | -101.101
5.3627 | -101.1017
5.3664 | -101.1012
5.3701 | -101.0993
5.3737 | -101.0962
5.3774 | -101.083
5.3811 | -101.0581
5.3848 | -101.019
5.3884 | -100.9704
5.3921 | -100.9142
5.3958 | -100.8661
5.3994 | -100.8208
5.4031 | -100.7833
5.4068 | -100.7535
5.4105 | -100.7398
5.4141 | -100.7406
5.4178 | -100.7525
5.4215 | -100.7736
5.4252 | -100.8027
5.4288 | -100.8376
5.4325 | -100.8708
5.4362 | -100.9057
5.4399 | -100.9461
5.4435 | -100.9758
5.4472 | -100.9903
5.4509 | -100.9962
5.4545 | -100.9994
5.4582 | -100.9927
5.4619 | -100.9812
5.4656 | -100.9664
5.4692 | -100.9517
5.4729 | -100.9411
5.4766 | -100.9263
5.4803 | -100.9055
5.4839 | -100.884
5.4876 | -100.874
5.4913 | -100.8748
5.4949 | -100.8852
5.4986 | -100.8984
5.5023 | -100.9151
5.506 | -100.9468
5.5096 | -100.9845
5.5133 | -101.0243
5.517 | -101.0647
5.5207 | -101.0978
5.5243 | -101.1262
5.528 | -101.159
5.5317 | -101.1852
5.5354 | -101.2075
5.539 | -101.2272
5.5427 | -101.2458
5.5464 | -101.2596
5.55 | -101.2763
5.5537 | -101.2903
5.5574 | -101.297
5.5611 | -101.3104
5.5647 | -101.3235
5.5684 | -101.3373
5.5721 | -101.351
5.5758 | -101.3676
5.5794 | -101.3933
5.5831 | -101.4311
5.5868 | -101.4733
5.5904 | -101.509
5.5941 | -101.5453
5.5978 | -101.5764
5.6015 | -101.5986
5.6051 | -101.6058
5.6088 | -101.5984
5.6125 | -101.5749
5.6162 | -101.5414
5.6198 | -101.4973
5.6235 | -101.4491
5.6272 | -101.4083
5.6309 | -101.3873
5.6345 | -101.3809
5.6382 | -101.3913
5.6419 | -101.4193
5.6455 | -101.4581
5.6492 | -101.5076
5.6529 | -101.5521
5.6566 | -101.5936
5.6602 | -101.6273
5.6639 | -101.6411
5.6676 | -101.6381
5.6713 | -101.6212
5.6749 | -101.5903
5.6786 | -101.557
5.6823 | -101.5326
5.686 | -101.5105
5.6896 | -101.4945
5.6933 | -101.4924
5.697 | -101.4969
5.7006 | -101.5134
5.7043 | -101.5328
5.708 | -101.5486
5.7117 | -101.5619
5.7153 | -101.5685
5.719 | -101.566
5.7227 | -101.5571
5.7264 | -101.5485
5.73 | -101.5361
5.7337 | -101.519
5.7374 | -101.5017
5.741 | -101.48
5.7447 | -101.4542
5.7484 | -101.4238
5.7521 | -101.3864
5.7557 | -101.347
5.7594 | -101.3075
5.7631 | -101.269
5.7668 | -101.2363
5.7704 | -101.2111
5.7741 | -101.192
5.7778 | -101.1794
5.7815 | -101.1715
5.7851 | -101.1598
5.7888 | -101.1448
5.7925 | -101.1311
5.7961 | -101.1169
5.7998 | -101.1031
5.8035 | -101.1026
5.8072 | -101.1198
5.8108 | -101.1406
5.8145 | -101.1786
5.8182 | -101.2336
5.8219 | -101.2898
5.8255 | -101.3299
5.8292 | -101.3684
5.8329 | -101.3995
5.8365 | -101.4218
5.8402 | -101.426
5.8439 | -101.415
5.8476 | -101.3958
5.8512 | -101.3664
5.8549 | -101.342
5.8586 | -101.3239
5.8623 | -101.3096
5.8659 | -101.2988
5.8696 | -101.2932
5.8733 | -101.2945
5.877 | -101.2949
5.8806 | -101.2899
5.8843 | -101.28
5.888 | -101.2711
5.8916 | -101.2659
5.8953 | -101.2636
5.899 | -101.2709
5.9027 | -101.2782
5.9063 | -101.2927
5.91 | -101.3217
5.9137 | -101.3551
5.9174 | -101.384
5.921 | -101.4072
5.9247 | -101.4224
5.9284 | -101.4305
5.932 | -101.4242
5.9357 | -101.4084
5.9394 | -101.3873
5.9431 | -101.3604
5.9467 | -101.3239
5.9504 | -101.2897
5.9541 | -101.2564
5.9578 | -101.218
5.9614 | -101.1783
5.9651 | -101.1406
5.9688 | -101.1079
5.9725 | -101.0787
5.9761 | -101.0545
5.9798 | -101.0375
5.9835 | -101.0283
5.9871 | -101.0369
5.9908 | -101.0627
5.9945 | -101.1053
5.9982 | -101.1545
6.0018 | -101.2116
6.0055 | -101.2801
6.0092 | -101.3615
6.0129 | -101.4452
6.0165 | -101.5308
6.0202 | -101.6094
6.0239 | -101.6881
6.0275 | -101.7554
6.0312 | -101.8116
6.0349 | -101.8603
6.0386 | -101.898
6.0422 | -101.9247
6.0459 | -101.9379
6.0496 | -101.9345
6.0533 | -101.9253
6.0569 | -101.9141
6.0606 | -101.8989
6.0643 | -101.8886
6.068 | -101.8826
6.0716 | -101.8785
6.0753 | -101.8721
6.079 | -101.863
6.0826 | -101.8541
6.0863 | -101.8434
6.09 | -101.8356
6.0937 | -101.8253
6.0973 | -101.8116
6.101 | -101.8029
6.1047 | -101.8049
6.1084 | -101.8164
6.112 | -101.8455
6.1157 | -101.8887
6.1194 | -101.9481
6.123 | -102.0119
6.1267 | -102.0939
6.1304 | -102.1887
6.1341 | -102.2821
6.1377 | -102.3784
6.1414 | -102.4655
6.1451 | -102.5239
6.1488 | -102.5542
6.1524 | -102.5631
6.1561 | -102.5548
6.1598 | -102.5341
6.1635 | -102.5007
6.1671 | -102.4606
6.1708 | -102.4151
6.1745 | -102.369
6.1781 | -102.3218
6.1818 | -102.2755
6.1855 | -102.2221
6.1892 | -102.1701
6.1928 | -102.1143
6.1965 | -102.0573
6.2002 | -102.0017
6.2039 | -101.9371
6.2075 | -101.87
6.2112 | -101.8108
6.2149 | -101.7622
6.2185 | -101.7117
6.2222 | -101.6661
6.2259 | -101.6287
6.2296 | -101.5989
6.2332 | -101.5823
6.2369 | -101.579
6.2406 | -101.5829
6.2443 | -101.5946
6.2479 | -101.6089
6.2516 | -101.6293
6.2553 | -101.6694
6.259 | -101.7141
6.2626 | -101.7575
6.2663 | -101.7984
6.27 | -101.8445
6.2736 | -101.8821
6.2773 | -101.918
6.281 | -101.9373
6.2847 | -101.945
6.2883 | -101.938
6.292 | -101.9159
6.2957 | -101.895
6.2994 | -101.8731
6.303 | -101.8456
6.3067 | -101.8276
6.3104 | -101.8188
6.314 | -101.8193
6.3177 | -101.8354
6.3214 | -101.8564
6.3251 | -101.884
6.3287 | -101.912
6.3324 | -101.941
6.3361 | -101.9788
6.3398 | -102.0164
6.3434 | -102.0521
6.3471 | -102.0769
6.3508 | -102.1037
6.3545 | -102.1342
6.3581 | -102.1466
6.3618 | -102.1623
6.3655 | -102.1709
6.3691 | -102.1785
6.3728 | -102.179
6.3765 | -102.1749
6.3802 | -102.1675
6.3838 | -102.1565
6.3875 | -102.1373
6.3912 | -102.1098
6.3949 | -102.078
6.3985 | -102.0484
6.4022 | -102.0199
6.4059 | -101.9975
6.4096 | -101.989
6.4132 | -101.9902
6.4169 | -102.0079
6.4206 | -102.0391
6.4242 | -102.0893
6.4279 | -102.1537
6.4316 | -102.2216
6.4353 | -102.3046
6.4389 | -102.3905
6.4426 | -102.4792
6.4463 | -102.5689
6.45 | -102.6464
6.4536 | -102.7121
6.4573 | -102.7718
6.461 | -102.8091
6.4646 | -102.8152
6.4683 | -102.8006
6.472 | -102.7704
6.4757 | -102.7223
6.4793 | -102.6577
6.483 | -102.5868
6.4867 | -102.5192
6.4904 | -102.4575
6.494 | -102.4023
6.4977 | -102.354
6.5014 | -102.3149
6.5051 | -102.289
6.5087 | -102.269
6.5124 | -102.2566
6.5161 | -102.254
6.5197 | -102.2587
6.5234 | -102.2663
6.5271 | -102.2773
6.5308 | -102.2957
6.5344 | -102.3097
6.5381 | -102.3303
6.5418 | -102.347
6.5455 | -102.3598
6.5491 | -102.3751
6.5528 | -102.3843
6.5565 | -102.3853
6.5601 | -102.3832
6.5638 | -102.3785
6.5675 | -102.3677
6.5712 | -102.3489
6.5748 | -102.3301
6.5785 | -102.3124
6.5822 | -102.2977
6.5859 | -102.2857
6.5895 | -102.2757
6.5932 | -102.2746
6.5969 | -102.2806
6.6006 | -102.2973
6.6042 | -102.3162
6.6079 | -102.3402
6.6116 | -102.3686
6.6152 | -102.3899
6.6189 | -102.4133
6.6226 | -102.4336
6.6263 | -102.443
6.6299 | -102.4436
6.6336 | -102.4404
6.6373 | -102.4265
6.641 | -102.4075
6.6446 | -102.3836
6.6483 | -102.3556
6.652 | -102.327
6.6556 | -102.2989
6.6593 | -102.2752
6.663 | -102.2563
6.6667 | -102.2437
6.6703 | -102.2438
6.674 | -102.2452
6.6777 | -102.2515
6.6814 | -102.2658
6.685 | -102.2813
6.6887 | -102.2906
6.6924 | -102.2978
6.6961 | -102.3024
6.6997 | -102.2994
6.7034 | -102.2912
6.7071 | -102.2718
6.7107 | -102.2483
6.7144 | -102.2287
6.7181 | -102.2023
6.7218 | -102.1803
6.7254 | -102.1655
6.7291 | -102.1567
6.7328 | -102.1549
6.7365 | -102.164
6.7401 | -102.1856
6.7438 | -102.2084
6.7475 | -102.2356
6.7511 | -102.269
6.7548 | -102.2996
6.7585 | -102.3263
6.7622 | -102.3472
6.7658 | -102.361
6.7695 | -102.3629
6.7732 | -102.3642
6.7769 | -102.3621
6.7805 | -102.3513
6.7842 | -102.3412
6.7879 | -102.3413
6.7916 | -102.3403
6.7952 | -102.3386
6.7989 | -102.3438
6.8026 | -102.3574
6.8062 | -102.3802
6.8099 | -102.4014
6.8136 | -102.4276
6.8173 | -102.4666
6.8209 | -102.5039
6.8246 | -102.5438
6.8283 | -102.5795
6.832 | -102.6027
6.8356 | -102.6219
6.8393 | -102.6364
6.843 | -102.6506
6.8466 | -102.6538
6.8503 | -102.65
6.854 | -102.6422
6.8577 | -102.631
6.8613 | -102.6181
6.865 | -102.6049
6.8687 | -102.5962
6.8724 | -102.5883
6.876 | -102.5828
6.8797 | -102.5811
6.8834 | -102.581
6.8871 | -102.5872
6.8907 | -102.5891
6.8944 | -102.5947
6.8981 | -102.6064
6.9017 | -102.6207
6.9054 | -102.639
6.9091 | -102.667
6.9128 | -102.7072
6.9164 | -102.7481
6.9201 | -102.7882
6.9238 | -102.8398
6.9275 | -102.8922
6.9311 | -102.9426
6.9348 | -102.9994
6.9385 | -103.0416
6.9421 | -103.0753
6.9458 | -103.1015
6.9495 | -103.1133
6.9532 | -103.115
6.9568 | -103.1017
6.9605 | -103.0807
6.9642 | -103.0535
6.9679 | -103.0166
6.9715 | -102.9717
6.9752 | -102.9237
6.9789 | -102.8756
6.9826 | -102.8319
6.9862 | -102.7914
6.9899 | -102.758
6.9936 | -102.7375
6.9972 | -102.7248
7.0009 | -102.7224
7.0046 | -102.7243
7.0083 | -102.7356
7.0119 | -102.755
7.0156 | -102.7745
7.0193 | -102.7887
7.023 | -102.8046
7.0266 | -102.8216
7.0303 | -102.8326
7.034 | -102.8389
7.0376 | -102.8388
7.0413 | -102.8295
7.045 | -102.8215
7.0487 | -102.8084
7.0523 | -102.7928
7.056 | -102.7817
7.0597 | -102.7732
7.0634 | -102.7679
7.067 | -102.7574
7.0707 | -102.7462
7.0744 | -102.7404
7.0781 | -102.7353
7.0817 | -102.7352
7.0854 | -102.7344
7.0891 | -102.7315
7.0927 | -102.7302
7.0964 | -102.7232
7.1001 | -102.715
7.1038 | -102.7102
7.1074 | -102.7039
7.1111 | -102.7047
7.1148 | -102.7066
7.1185 | -102.7112
7.1221 | -102.7253
7.1258 | -102.7488
7.1295 | -102.7695
7.1331 | -102.7927
7.1368 | -102.8152
7.1405 | -102.838
7.1442 | -102.8665
7.1478 | -102.8893
7.1515 | -102.9007
7.1552 | -102.9059
7.1589 | -102.9066
7.1625 | -102.9031
7.1662 | -102.8922
7.1699 | -102.8731
7.1736 | -102.8507
7.1772 | -102.8275
7.1809 | -102.7991
7.1846 | -102.77
7.1882 | -102.7386
7.1919 | -102.7137
7.1956 | -102.6946
7.1993 | -102.6802
7.2029 | -102.6672
7.2066 | -102.6553
7.2103 | -102.6453
7.214 | -102.6435
7.2176 | -102.6418
7.2213 | -102.6429
7.225 | -102.6485
7.2287 | -102.6534
7.2323 | -102.6664
7.236 | -102.6906
7.2397 | -102.7163
7.2433 | -102.748
7.247 | -102.7876
7.2507 | -102.8295
7.2544 | -102.8756
7.258 | -102.9264
7.2617 | -102.9795
7.2654 | -103.0293
7.2691 | -103.0726
7.2727 | -103.1042
7.2764 | -103.1329
7.2801 | -103.1522
7.2837 | -103.1559
7.2874 | -103.1523
7.2911 | -103.1368
7.2948 | -103.1157
7.2984 | -103.0895
7.3021 | -103.0597
7.3058 | -103.0283
7.3095 | -103.0002
7.3131 | -102.9742
7.3168 | -102.9524
7.3205 | -102.9297
7.3242 | -102.913
7.3278 | -102.9028
7.3315 | -102.9016
7.3352 | -102.909
7.3388 | -102.9208
7.3425 | -102.9425
7.3462 | -102.9773
7.3499 | -103.0175
7.3535 | -103.0601
7.3572 | -103.1063
7.3609 | -103.1576
7.3646 | -103.2115
7.3682 | -103.2593
7.3719 | -103.3041
7.3756 | -103.3381
7.3792 | -103.3652
7.3829 | -103.3769
7.3866 | -103.3767
7.3903 | -103.3744
7.3939 | -103.3593
7.3976 | -103.3338
7.4013 | -103.3054
7.405 | -103.2743
7.4086 | -103.2423
7.4123 | -103.2136
7.416 | -103.1932
7.4197 | -103.1759
7.4233 | -103.1566
7.427 | -103.1402
7.4307 | -103.1311
7.4343 | -103.1293
7.438 | -103.1317
7.4417 | -103.1351
7.4454 | -103.1484
7.449 | -103.1675
7.4527 | -103.1871
7.4564 | -103.2099
7.4601 | -103.2421
7.4637 | -103.2781
7.4674 | -103.3119
7.4711 | -103.3435
7.4747 | -103.3726
7.4784 | -103.4018
7.4821 | -103.4304
7.4858 | -103.4593
7.4894 | -103.4827
7.4931 | -103.5027
7.4968 | -103.5171
7.5005 | -103.5208
7.5041 | -103.5194
7.5078 | -103.5232
7.5115 | -103.5173
7.5152 | -103.5091
7.5188 | -103.5025
7.5225 | -103.4928
7.5262 | -103.493
7.5298 | -103.4977
7.5335 | -103.4994
7.5372 | -103.5075
7.5409 | -103.5142
7.5445 | -103.5181
7.5482 | -103.5176
7.5519 | -103.51
7.5556 | -103.5031
7.5592 | -103.4917
7.5629 | -103.4803
7.5666 | -103.4693
7.5702 | -103.4571
7.5739 | -103.446
7.5776 | -103.4384
7.5813 | -103.4351
7.5849 | -103.4316
7.5886 | -103.4335
7.5923 | -103.4401
7.596 | -103.4486
7.5996 | -103.4658
7.6033 | -103.4866
7.607 | -103.5048
7.6107 | -103.5205
7.6143 | -103.5357
7.618 | -103.5492
7.6217 | -103.5529
7.6253 | -103.5542
7.629 | -103.5519
7.6327 | -103.5424
7.6364 | -103.5267
7.64 | -103.5044
7.6437 | -103.4868
7.6474 | -103.4731
7.6511 | -103.4629
7.6547 | -103.4605
7.6584 | -103.4651
7.6621 | -103.4766
7.6657 | -103.4943
7.6694 | -103.51
7.6731 | -103.5287
7.6768 | -103.5519
7.6804 | -103.5748
7.6841 | -103.5955
7.6878 | -103.6142
7.6915 | -103.6326
7.6951 | -103.6487
7.6988 | -103.6604
7.7025 | -103.6665
7.7062 | -103.665
7.7098 | -103.6616
7.7135 | -103.6545
7.7172 | -103.6488
7.7208 | -103.6474
7.7245 | -103.6479
7.7282 | -103.6468
7.7319 | -103.6495
7.7355 | -103.657
7.7392 | -103.6684
7.7429 | -103.6818
7.7466 | -103.7062
7.7502 | -103.7384
7.7539 | -103.766
7.7576 | -103.7914
7.7612 | -103.822
7.7649 | -103.8501
7.7686 | -103.8724
7.7723 | -103.8912
7.7759 | -103.9075
7.7796 | -103.9147
7.7833 | -103.9099
7.787 | -103.8982
7.7906 | -103.8845
7.7943 | -103.8647
7.798 | -103.8451
7.8017 | -103.8242
7.8053 | -103.8
7.809 | -103.7821
7.8127 | -103.7698
7.8163 | -103.7592
7.82 | -103.7479
7.8237 | -103.7376
7.8274 | -103.7291
7.831 | -103.7172
7.8347 | -103.7035
7.8384 | -103.6902
7.8421 | -103.6756
7.8457 | -103.6627
7.8494 | -103.6495
7.8531 | -103.6303
7.8567 | -103.6108
7.8604 | -103.5897
7.8641 | -103.5642
7.8678 | -103.5411
7.8714 | -103.5186
7.8751 | -103.4903
7.8788 | -103.464
7.8825 | -103.4408
7.8861 | -103.4225
7.8898 | -103.4049
7.8935 | -103.3933
7.8972 | -103.3942
7.9008 | -103.3953
7.9045 | -103.3965
7.9082 | -103.4092
7.9118 | -103.4263
7.9155 | -103.4475
7.9192 | -103.4711
7.9229 | -103.494
7.9265 | -103.5176
7.9302 | -103.5406
7.9339 | -103.5695
7.9376 | -103.5955
7.9412 | -103.6223
7.9449 | -103.649
7.9486 | -103.6586
7.9523 | -103.678
7.9559 | -103.6924
7.9596 | -103.7041
7.9633 | -103.7068
7.9669 | -103.7024
7.9706 | -103.6965
7.9743 | -103.6894
7.978 | -103.6748
7.9816 | -103.6623
7.9853 | -103.6428
7.989 | -103.6216
7.9927 | -103.6037
7.9963 | -103.5788
8 | -103.5544
8.0037 | -103.5248
8.0073 | -103.4975
8.011 | -103.4753
8.0147 | -103.4461
8.0184 | -103.42
8.022 | -103.3936
8.0257 | -103.3652
8.0294 | -103.3373
8.0331 | -103.3082
8.0367 | -103.2791
8.0404 | -103.2502
8.0441 | -103.2253
8.0478 | -103.1997
8.0514 | -103.176
8.0551 | -103.1574
8.0588 | -103.1371
8.0624 | -103.1225
8.0661 | -103.1168
8.0698 | -103.1154
8.0735 | -103.1152
8.0771 | -103.1204
8.0808 | -103.1313
8.0845 | -103.1457
8.0882 | -103.1646
8.0918 | -103.1895
8.0955 | -103.2161
8.0992 | -103.2494
8.1028 | -103.2839
8.1065 | -103.3252
8.1102 | -103.3669
8.1139 | -103.4098
8.1175 | -103.4595
8.1212 | -103.5086
8.1249 | -103.5581
8.1286 | -103.6049
8.1322 | -103.6485
8.1359 | -103.6937
8.1396 | -103.7342
8.1433 | -103.7722
8.1469 | -103.8081
8.1506 | -103.8369
8.1543 | -103.8608
8.1579 | -103.8792
8.1616 | -103.8943
8.1653 | -103.9023
8.169 | -103.9065
8.1726 | -103.9033
8.1763 | -103.894
8.18 | -103.8826
8.1837 | -103.8699
8.1873 | -103.8521
8.191 | -103.8389
8.1947 | -103.8287
8.1983 | -103.8179
8.202 | -103.8093
8.2057 | -103.8125
8.2094 | -103.819
8.213 | -103.8247
8.2167 | -103.831
8.2204 | -103.8449
8.2241 | -103.8631
8.2277 | -103.878
8.2314 | -103.8962
8.2351 | -103.9088
8.2388 | -103.919
8.2424 | -103.9301
8.2461 | -103.9406
8.2498 | -103.9469
8.2534 | -103.9515
8.2571 | -103.9493
8.2608 | -103.9496
8.2645 | -103.9491
8.2681 | -103.9431
8.2718 | -103.9349
8.2755 | -103.9302
8.2792 | -103.922
8.2828 | -103.9095
8.2865 | -103.8996
8.2902 | -103.8872
8.2938 | -103.8759
8.2975 | -103.8716
8.3012 | -103.8699
8.3049 | -103.8677
8.3085 | -103.8658
8.3122 | -103.8639
8.3159 | -103.8688
8.3196 | -103.8771
8.3232 | -103.8849
8.3269 | -103.8865
8.3306 | -103.891
8.3343 | -103.8977
8.3379 | -103.9065
8.3416 | -103.91
8.3453 | -103.9116
8.3489 | -103.918
8.3526 | -103.9225
8.3563 | -103.936
8.36 | -103.9489
8.3636 | -103.9618
8.3673 | -103.9809
8.371 | -104.0028
8.3747 | -104.0273
8.3783 | -104.0568
8.382 | -104.0883
8.3857 | -104.1176
8.3893 | -104.15
8.393 | -104.1812
8.3967 | -104.1996
8.4004 | -104.2159
8.404 | -104.2252
8.4077 | -104.2256
8.4114 | -104.2229
8.4151 | -104.2105
8.4187 | -104.1843
8.4224 | -104.1544
8.4261 | -104.1192
8.4298 | -104.0852
8.4334 | -104.0518
8.4371 | -104.0165
8.4408 | -103.985
8.4444 | -103.9612
8.4481 | -103.9428
8.4518 | -103.9284
8.4555 | -103.9208
8.4591 | -103.9156
8.4628 | -103.9168
8.4665 | -103.9261
8.4702 | -103.9352
8.4738 | -103.9476
8.4775 | -103.9677
8.4812 | -103.9781
8.4848 | -103.9957
8.4885 | -104.0161
8.4922 | -104.0378
8.4959 | -104.0648
8.4995 | -104.0876
8.5032 | -104.1105
8.5069 | -104.1347
8.5106 | -104.1614
8.5142 | -104.1863
8.5179 | -104.2085
8.5216 | -104.2307
8.5253 | -104.2483
8.5289 | -104.2574
8.5326 | -104.2595
8.5363 | -104.2584
8.5399 | -104.2523
8.5436 | -104.2363
8.5473 | -104.2142
8.551 | -104.1912
8.5546 | -104.1596
8.5583 | -104.1213
8.562 | -104.0824
8.5657 | -104.0456
8.5693 | -104.0081
8.573 | -103.9759
8.5767 | -103.9459
8.5803 | -103.9197
8.584 | -103.8993
8.5877 | -103.8872
8.5914 | -103.8825
8.595 | -103.8848
8.5987 | -103.8904
8.6024 | -103.9005
8.6061 | -103.9175
8.6097 | -103.9397
8.6134 | -103.9599
8.6171 | -103.9821
8.6208 | -104.009
8.6244 | -104.037
8.6281 | -104.067
8.6318 | -104.0975
8.6354 | -104.1201
8.6391 | -104.1393
8.6428 | -104.1608
8.6465 | -104.185
8.6501 | -104.2059
8.6538 | -104.2249
8.6575 | -104.2448
8.6612 | -104.263
8.6648 | -104.2794
8.6685 | -104.2919
8.6722 | -104.3055
8.6758 | -104.3168
8.6795 | -104.3216
8.6832 | -104.3252
8.6869 | -104.3297
8.6905 | -104.3349
8.6942 | -104.337
8.6979 | -104.3384
8.7016 | -104.3396
8.7052 | -104.3373
8.7089 | -104.3266
8.7126 | -104.3173
8.7163 | -104.3031
8.7199 | -104.2826
8.7236 | -104.2599
8.7273 | -104.2376
8.7309 | -104.2131
8.7346 | -104.1862
8.7383 | -104.1572
8.742 | -104.1292
8.7456 | -104.1004
8.7493 | -104.0677
8.753 | -104.0384
8.7567 | -104.0146
8.7603 | -103.9912
8.764 | -103.9726
8.7677 | -103.9587
8.7713 | -103.9418
8.775 | -103.9306
8.7787 | -103.9191
8.7824 | -103.9112
8.786 | -103.9092
8.7897 | -103.9137
8.7934 | -103.9215
8.7971 | -103.9297
8.8007 | -103.9372
8.8044 | -103.9431
8.8081 | -103.9511
8.8118 | -103.9557
8.8154 | -103.9564
8.8191 | -103.9511
8.8228 | -103.9407
8.8264 | -103.9295
8.8301 | -103.917
8.8338 | -103.9065
8.8375 | -103.8947
8.8411 | -103.8775
8.8448 | -103.8565
8.8485 | -103.8395
8.8522 | -103.8281
8.8558 | -103.8169
8.8595 | -103.8097
8.8632 | -103.8068
8.8668 | -103.8089
8.8705 | -103.8177
8.8742 | -103.8323
8.8779 | -103.8523
8.8815 | -103.8698
8.8852 | -103.8887
8.8889 | -103.9108
8.8926 | -103.9317
8.8962 | -103.9507
8.8999 | -103.9663
8.9036 | -103.974
8.9073 | -103.9869
8.9109 | -103.9903
8.9146 | -103.9878
8.9183 | -103.9823
8.9219 | -103.9775
8.9256 | -103.9766
8.9293 | -103.9764
8.933 | -103.973
8.9366 | -103.9685
8.9403 | -103.9709
8.944 | -103.9796
8.9477 | -103.9941
8.9513 | -104.0158
8.955 | -104.0416
8.9587 | -104.0747
8.9624 | -104.1132
8.966 | -104.1615
8.9697 | -104.2136
8.9734 | -104.2737
8.977 | -104.3342
8.9807 | -104.3909
8.9844 | -104.4428
8.9881 | -104.4908
8.9917 | -104.5337
8.9954 | -104.5732
8.9991 | -104.6031
9.0028 | -104.6199
9.0064 | -104.5993
9.0101 | -104.6005
9.0138 | -104.5959
9.0174 | -104.5827
9.0211 | -104.563
9.0248 | -104.5397
9.0285 | -104.5097
9.0321 | -104.4744
9.0358 | -104.4378
9.0395 | -104.4003
9.0432 | -104.3689
9.0468 | -104.3333
9.0505 | -104.2967
9.0542 | -104.2653
9.0579 | -104.2398
9.0615 | -104.2173
9.0652 | -104.2019
9.0689 | -104.1922
9.0725 | -104.1867
9.0762 | -104.1826
9.0799 | -104.1827
9.0836 | -104.1848
9.0872 | -104.1893
9.0909 | -104.1967
9.0946 | -104.2073
9.0983 | -104.2114
9.1019 | -104.2151
9.1056 | -104.2162
9.1093 | -104.2158
9.1129 | -104.2154
9.1166 | -104.2141
9.1203 | -104.204
9.124 | -104.194
9.1276 | -104.1781
9.1313 | -104.1578
9.135 | -104.1422
9.1387 | -104.1273
9.1423 | -104.1118
9.146 | -104.0962
9.1497 | -104.0833
9.1534 | -104.0731
9.157 | -104.0676
9.1607 | -104.0667
9.1644 | -104.0729
9.168 | -104.0853
9.1717 | -104.1037
9.1754 | -104.1248
9.1791 | -104.1515
9.1827 | -104.1875
9.1864 | -104.231
9.1901 | -104.2777
9.1938 | -104.3235
9.1974 | -104.3738
9.2011 | -104.4264
9.2048 | -104.4726
9.2084 | -104.5153
9.2121 | -104.557
9.2158 | -104.5937
9.2195 | -104.6303
9.2231 | -104.6621
9.2268 | -104.6866
9.2305 | -104.7023
9.2342 | -104.7156
9.2378 | -104.7182
9.2415 | -104.7156
9.2452 | -104.7062
9.2489 | -104.6906
9.2525 | -104.6727
9.2562 | -104.6548
9.2599 | -104.6293
9.2635 | -104.6016
9.2672 | -104.5746
9.2709 | -104.5437
9.2746 | -104.5091
9.2782 | -104.4791
9.2819 | -104.4517
9.2856 | -104.4257
9.2893 | -104.4001
9.2929 | -104.3759
9.2966 | -104.3573
9.3003 | -104.3434
9.3039 | -104.3344
9.3076 | -104.3254
9.3113 | -104.322
9.315 | -104.3246
9.3186 | -104.334
9.3223 | -104.3417
9.326 | -104.3494
9.3297 | -104.3626
9.3333 | -104.3766
9.337 | -104.3915
9.3407 | -104.4093
9.3444 | -104.4254
9.348 | -104.4445
9.3517 | -104.4611
9.3554 | -104.4681
9.359 | -104.4737
9.3627 | -104.4761
9.3664 | -104.4735
9.3701 | -104.4707
9.3737 | -104.467
9.3774 | -104.4589
9.3811 | -104.446
9.3848 | -104.4312
9.3884 | -104.4172
9.3921 | -104.4001
9.3958 | -104.3857
9.3994 | -104.3746
9.4031 | -104.3638
9.4068 | -104.3573
9.4105 | -104.3568
9.4141 | -104.3547
9.4178 | -104.3604
9.4215 | -104.3748
9.4252 | -104.3941
9.4288 | -104.4145
9.4325 | -104.4428
9.4362 | -104.4759
9.4399 | -104.5109
9.4435 | -104.5503
9.4472 | -104.589
9.4509 | -104.623
9.4545 | -104.6575
9.4582 | -104.6934
9.4619 | -104.7277
9.4656 | -104.7577
9.4692 | -104.7769
9.4729 | -104.791
9.4766 | -104.803
9.4803 | -104.8106
9.4839 | -104.81
9.4876 | -104.8107
9.4913 | -104.8133
9.4949 | -104.8131
9.4986 | -104.8134
9.5023 | -104.8158
9.506 | -104.81
9.5096 | -104.8049
9.5133 | -104.7995
9.517 | -104.7985
9.5207 | -104.7944
9.5243 | -104.788
9.528 | -104.7836
9.5317 | -104.7785
9.5354 | -104.7725
9.539 | -104.7663
9.5427 | -104.7592
9.5464 | -104.7532
9.55 | -104.745
9.5537 | -104.7362
9.5574 | -104.7246
9.5611 | -104.7108
9.5647 | -104.6971
9.5684 | -104.6834
9.5721 | -104.6652
9.5758 | -104.6446
9.5794 | -104.6251
9.5831 | -104.602
9.5868 | -104.5813
9.5905 | -104.5622
9.5941 | -104.5426
9.5978 | -104.5277
9.6015 | -104.5164
9.6051 | -104.5109
9.6088 | -104.5157
9.6125 | -104.5206
9.6162 | -104.5306
9.6198 | -104.5414
9.6235 | -104.5604
9.6272 | -104.5868
9.6309 | -104.6179
9.6345 | -104.6495
9.6382 | -104.6794
9.6419 | -104.7147
9.6455 | -104.7507
9.6492 | -104.781
9.6529 | -104.81
9.6566 | -104.835
9.6602 | -104.8559
9.6639 | -104.8682
9.6676 | -104.8791
9.6713 | -104.8903
9.6749 | -104.8992
9.6786 | -104.9007
9.6823 | -104.8934
9.686 | -104.8839
9.6896 | -104.8678
9.6933 | -104.8475
9.697 | -104.8203
9.7006 | -104.787
9.7043 | -104.7488
9.708 | -104.7094
9.7117 | -104.6705
9.7153 | -104.6244
9.719 | -104.5728
9.7227 | -104.5236
9.7264 | -104.4763
9.73 | -104.4277
9.7337 | -104.3822
9.7374 | -104.3423
9.741 | -104.3067
9.7447 | -104.2762
9.7484 | -104.2486
9.7521 | -104.2259
9.7557 | -104.2069
9.7594 | -104.194
9.7631 | -104.1911
9.7668 | -104.1908
9.7704 | -104.1906
9.7741 | -104.1967
9.7778 | -104.2124
9.7815 | -104.235
9.7851 | -104.2567
9.7888 | -104.2813
9.7925 | -104.3108
9.7961 | -104.3459
9.7998 | -104.3822
9.8035 | -104.4218
9.8072 | -104.4654
9.8108 | -104.5046
9.8145 | -104.5356
9.8182 | -104.5643
9.8219 | -104.5976
9.8255 | -104.6294
9.8292 | -104.6581
9.8329 | -104.6849
9.8365 | -104.7111
9.8402 | -104.7386
9.8439 | -104.7645
9.8476 | -104.7856
9.8512 | -104.8054
9.8549 | -104.8231
9.8586 | -104.8444
9.8623 | -104.8734
9.8659 | -104.8985
9.8696 | -104.9217
9.8733 | -104.9493
9.877 | -104.9771
9.8806 | -104.9985
9.8843 | -105.0243
9.888 | -105.0527
9.8916 | -105.0805
9.8953 | -105.1087
9.899 | -105.142
9.9027 | -105.1689
9.9063 | -105.1925
9.91 | -105.2117
9.9137 | -105.2268
9.9174 | -105.2343
9.921 | -105.2377
9.9247 | -105.241
9.9284 | -105.2391
9.932 | -105.2358
9.9357 | -105.2245
9.9394 | -105.2018
9.9431 | -105.1763
9.9467 | -105.148
9.9504 | -105.1185
9.9541 | -105.0919
9.9578 | -105.0707
9.9614 | -105.0505
9.9651 | -105.0291
9.9688 | -105.0101
9.9725 | -104.9944
9.9761 | -104.9789
9.9798 | -104.9687
9.9835 | -104.9576
9.9871 | -104.9525
9.9908 | -104.9525
9.9945 | -104.9545
9.9982 | -104.9596
10.0018 | -104.9664
10.0055 | -104.9683
10.0092 | -104.9708
10.0129 | -104.975
10.0165 | -104.9744
10.0202 | -104.9735
10.0239 | -104.9753
10.0275 | -104.9723
10.0312 | -104.9683
10.0349 | -104.9672
10.0386 | -104.9638
10.0422 | -104.9608
10.0459 | -104.9613
10.0496 | -104.9651
10.0533 | -104.967
10.0569 | -104.9734
10.0606 | -104.9837
10.0643 | -104.993
10.068 | -105.0148
10.0716 | -105.0274
10.0753 | -105.0466
10.079 | -105.0628
10.0826 | -105.0783
10.0863 | -105.0959
10.09 | -105.1162
10.0937 | -105.1302
10.0973 | -105.1408
10.101 | -105.1514
10.1047 | -105.1594
10.1084 | -105.1664
10.112 | -105.1696
10.1157 | -105.1719
10.1194 | -105.171
10.123 | -105.1693
10.1267 | -105.1629
10.1304 | -105.1533
10.1341 | -105.1379
10.1377 | -105.1243
10.1414 | -105.1107
10.1451 | -105.0962
10.1488 | -105.0863
10.1524 | -105.0801
10.1561 | -105.0749
10.1598 | -105.0712
10.1635 | -105.0685
10.1671 | -105.0672
10.1708 | -105.0678
10.1745 | -105.0703
10.1781 | -105.0738
10.1818 | -105.0799
10.1855 | -105.0896
10.1892 | -105.0989
10.1928 | -105.1094
10.1965 | -105.1204
10.2002 | -105.136
10.2039 | -105.1536
10.2075 | -105.1662
10.2112 | -105.1716
10.2149 | -105.1744
10.2185 | -105.1774
10.2222 | -105.1763
10.2259 | -105.1755
10.2296 | -105.1722
10.2332 | -105.165
10.2369 | -105.1535
10.2406 | -105.136
10.2443 | -105.114
10.2479 | -105.0873
10.2516 | -105.0577
10.2553 | -105.0291
10.259 | -105.0024
10.2626 | -104.9784
10.2663 | -104.9542
10.27 | -104.9293
10.2736 | -104.9068
10.2773 | -104.8857
10.281 | -104.8692
10.2847 | -104.8567
10.2883 | -104.8425
10.292 | -104.8317
10.2957 | -104.8203
10.2994 | -104.8061
10.303 | -104.8001
10.3067 | -104.7952
10.3104 | -104.7916
10.314 | -104.7909
10.3177 | -104.7891
10.3214 | -104.7888
10.3251 | -104.7895
10.3287 | -104.7882
10.3324 | -104.7876
10.3361 | -104.7868
10.3398 | -104.7874
10.3434 | -104.7913
10.3471 | -104.7945
10.3508 | -104.7964
10.3545 | -104.7965
10.3581 | -104.8004
10.3618 | -104.8081
10.3655 | -104.8167
10.3691 | -104.825
10.3728 | -104.8355
10.3765 | -104.8435
10.3802 | -104.8576
10.3838 | -104.8744
10.3875 | -104.8914
10.3912 | -104.9096
10.3949 | -104.9293
10.3985 | -104.9502
10.4022 | -104.9702
10.4059 | -104.9915
10.4095 | -105.0114
10.4132 | -105.0326
10.4169 | -105.0508
10.4206 | -105.0735
10.4242 | -105.0973
10.4279 | -105.1195
10.4316 | -105.1425
10.4353 | -105.1624
10.4389 | -105.1828
10.4426 | -105.2048
10.4463 | -105.2314
10.45 | -105.2565
10.4536 | -105.2756
10.4573 | -105.2946
10.461 | -105.3088
10.4646 | -105.3209
10.4683 | -105.3329
10.472 | -105.3409
10.4757 | -105.3476
10.4793 | -105.3543
10.483 | -105.3566
10.4867 | -105.3539
10.4904 | -105.3503
10.494 | -105.3443
10.4977 | -105.3395
10.5014 | -105.3336
10.5051 | -105.325
10.5087 | -105.3143
10.5124 | -105.3042
10.5161 | -105.2939
10.5197 | -105.2829
10.5234 | -105.2698
10.5271 | -105.258
10.5308 | -105.244
10.5344 | -105.232
10.5381 | -105.221
10.5418 | -105.2105
10.5455 | -105.204
10.5491 | -105.1914
10.5528 | -105.179
10.5565 | -105.1673
10.5601 | -105.16
10.5638 | -105.1532
10.5675 | -105.153
10.5712 | -105.1493
10.5748 | -105.1464
10.5785 | -105.1498
10.5822 | -105.1518
10.5859 | -105.1549
10.5895 | -105.1603
10.5932 | -105.164
10.5969 | -105.1695
10.6006 | -105.1729
10.6042 | -105.1788
10.6079 | -105.1846
10.6116 | -105.1946
10.6152 | -105.2026
10.6189 | -105.2083
10.6226 | -105.2146
10.6263 | -105.2228
10.6299 | -105.2303
10.6336 | -105.2352
10.6373 | -105.2351
10.641 | -105.2373
10.6446 | -105.2381
10.6483 | -105.236
10.652 | -105.232
10.6556 | -105.23
10.6593 | -105.2271
10.663 | -105.2193
10.6667 | -105.2048
10.6703 | -105.1901
10.674 | -105.1741
10.6777 | -105.16
10.6814 | -105.1479
10.685 | -105.1343
10.6887 | -105.1207
10.6924 | -105.1064
10.6961 | -105.0907
10.6997 | -105.0774
10.7034 | -105.0671
10.7071 | -105.0572
10.7107 | -105.0496
10.7144 | -105.0439
10.7181 | -105.0409
10.7218 | -105.0374
10.7254 | -105.0372
10.7291 | -105.0373
10.7328 | -105.0396
10.7365 | -105.0466
10.7401 | -105.0529
10.7438 | -105.0595
10.7475 | -105.069
10.7511 | -105.0829
10.7548 | -105.0936
10.7585 | -105.0999
10.7622 | -105.1076
10.7658 | -105.1145
10.7695 | -105.1182
10.7732 | -105.1246
10.7769 | -105.1289
10.7805 | -105.1269
10.7842 | -105.1208
10.7879 | -105.1148
10.7916 | -105.1138
10.7952 | -105.1108
10.7989 | -105.1028
10.8026 | -105.0965
10.8062 | -105.0923
10.8099 | -105.088
10.8136 | -105.0838
10.8173 | -105.0746
10.8209 | -105.0681
10.8246 | -105.0662
10.8283 | -105.0628
10.832 | -105.0597
10.8356 | -105.0589
10.8393 | -105.0586
10.843 | -105.0595
10.8466 | -105.0614
10.8503 | -105.0618
10.854 | -105.0628
10.8577 | -105.0624
10.8613 | -105.0645
10.865 | -105.0652
10.8687 | -105.0616
10.8724 | -105.054
10.876 | -105.0481
10.8797 | -105.0445
10.8834 | -105.0383
10.8871 | -105.0318
10.8907 | -105.0235
10.8944 | -105.0139
10.8981 | -105.0084
10.9017 | -105.0015
10.9054 | -104.9912
10.9091 | -104.9828
10.9128 | -104.9779
10.9164 | -104.9766
10.9201 | -104.9751
10.9238 | -104.9722
10.9275 | -104.9719
10.9311 | -104.9781
10.9348 | -104.9873
10.9385 | -105.002
10.9421 | -105.0221
10.9458 | -105.043
10.9495 | -105.062
10.9532 | -105.0868
10.9568 | -105.1109
10.9605 | -105.1336
10.9642 | -105.1575
10.9679 | -105.1865
10.9715 | -105.2132
10.9752 | -105.2441
10.9789 | -105.2719
10.9826 | -105.2929
10.9862 | -105.311
10.9899 | -105.3254
10.9936 | -105.3345
10.9972 | -105.3415
11.0009 | -105.3492
11.0046 | -105.3557
11.0083 | -105.3638
11.0119 | -105.3692
11.0156 | -105.3698
11.0193 | -105.3717
11.023 | -105.3745
11.0266 | -105.379
11.0303 | -105.3849
11.034 | -105.3878
11.0376 | -105.3884
11.0413 | -105.3899
11.045 | -105.3893
11.0487 | -105.3875
11.0523 | -105.386
11.056 | -105.3809
11.0597 | -105.3756
11.0634 | -105.369
11.067 | -105.3593
11.0707 | -105.3503
11.0744 | -105.3393
11.0781 | -105.3298
11.0817 | -105.32
11.0854 | -105.305
11.0891 | -105.2853
11.0927 | -105.2676
11.0964 | -105.2516
11.1001 | -105.2311
11.1038 | -105.2141
11.1074 | -105.2035
11.1111 | -105.1929
11.1148 | -105.1833
11.1185 | -105.1718
11.1221 | -105.1636
11.1258 | -105.1659
11.1295 | -105.1591
11.1332 | -105.1593
11.1368 | -105.159
11.1405 | -105.1586
11.1442 | -105.1629
11.1478 | -105.1666
11.1515 | -105.1747
11.1552 | -105.1845
11.1589 | -105.1901
11.1625 | -105.1987
11.1662 | -105.2073
11.1699 | -105.2177
11.1736 | -105.2291
11.1772 | -105.2415
11.1809 | -105.2474
11.1846 | -105.2525
11.1882 | -105.2609
11.1919 | -105.2662
11.1956 | -105.2662
11.1993 | -105.2697
11.2029 | -105.2704
11.2066 | -105.2703
11.2103 | -105.2693
11.214 | -105.2621
11.2176 | -105.256
11.2213 | -105.2471
11.225 | -105.2321
11.2287 | -105.2154
11.2323 | -105.1971
11.236 | -105.1748
11.2397 | -105.148
11.2433 | -105.1228
11.247 | -105.097
11.2507 | -105.0713
11.2544 | -105.046
11.258 | -105.0199
11.2617 | -104.9983
11.2654 | -104.9751
11.2691 | -104.9514
11.2727 | -104.9311
11.2764 | -104.9127
11.2801 | -104.8968
11.2837 | -104.8781
11.2874 | -104.8633
11.2911 | -104.8537
11.2948 | -104.8441
11.2984 | -104.8351
11.3021 | -104.8323
11.3058 | -104.8341
11.3095 | -104.8372
11.3131 | -104.8419
11.3168 | -104.8459
11.3205 | -104.851
11.3242 | -104.8578
11.3278 | -104.8633
11.3315 | -104.8675
11.3352 | -104.8704
11.3388 | -104.8743
11.3425 | -104.8797
11.3462 | -104.8865
11.3499 | -104.8948
11.3535 | -104.9036
11.3572 | -104.9142
11.3609 | -104.9297
11.3646 | -104.943
11.3682 | -104.9576
11.3719 | -104.9704
11.3756 | -104.9859
11.3792 | -105.0078
11.3829 | -105.0288
11.3866 | -105.0511
11.3903 | -105.0769
11.3939 | -105.1032
11.3976 | -105.13
11.4013 | -105.1616
11.405 | -105.1944
11.4086 | -105.2266
11.4123 | -105.2567
11.416 | -105.2849
11.4197 | -105.3111
11.4233 | -105.337
11.427 | -105.3589
11.4307 | -105.3793
11.4343 | -105.3971
11.438 | -105.4105
11.4417 | -105.4234
11.4454 | -105.4334
11.449 | -105.4419
11.4527 | -105.4461
11.4564 | -105.4462
11.4601 | -105.4419
11.4637 | -105.4347
11.4674 | -105.4246
11.4711 | -105.4147
11.4747 | -105.4044
11.4784 | -105.3964
11.4821 | -105.387
11.4858 | -105.3777
11.4894 | -105.3732
11.4931 | -105.3677
11.4968 | -105.3654
11.5005 | -105.3634
11.5041 | -105.3608
11.5078 | -105.3604
11.5115 | -105.3607
11.5152 | -105.3592
11.5188 | -105.3589
11.5225 | -105.3637
11.5262 | -105.3693
11.5298 | -105.3743
11.5335 | -105.3823
11.5372 | -105.3919
11.5409 | -105.3987
11.5445 | -105.4044
11.5482 | -105.4099
11.5519 | -105.4163
11.5556 | -105.4215
11.5592 | -105.4199
11.5629 | -105.4183
11.5666 | -105.4128
11.5702 | -105.4075
11.5739 | -105.4027
11.5776 | -105.3995
11.5813 | -105.3963
11.5849 | -105.3884
11.5886 | -105.3793
11.5923 | -105.3695
11.596 | -105.3591
11.5996 | -105.3512
11.6033 | -105.343
11.607 | -105.3358
11.6107 | -105.3276
11.6143 | -105.3204
11.618 | -105.3102
11.6217 | -105.3001
11.6253 | -105.299
11.629 | -105.2985
11.6327 | -105.2945
11.6364 | -105.2919
11.64 | -105.2884
11.6437 | -105.2822
11.6474 | -105.2779
11.6511 | -105.2727
11.6547 | -105.2621
11.6584 | -105.2551
11.6621 | -105.2472
11.6657 | -105.2382
11.6694 | -105.2317
11.6731 | -105.2256
11.6768 | -105.2169
11.6804 | -105.21
11.6841 | -105.2025
11.6878 | -105.1935
11.6915 | -105.1849
11.6951 | -105.1782
11.6988 | -105.1749
11.7025 | -105.171
11.7062 | -105.1722
11.7098 | -105.1744
11.7135 | -105.1754
11.7172 | -105.1797
11.7208 | -105.1865
11.7245 | -105.1946
11.7282 | -105.2024
11.7319 | -105.2098
11.7355 | -105.2199
11.7392 | -105.2326
11.7429 | -105.244
11.7466 | -105.257
11.7502 | -105.2686
11.7539 | -105.2817
11.7576 | -105.2951
11.7612 | -105.3075
11.7649 | -105.3217
11.7686 | -105.3335
11.7723 | -105.3441
11.7759 | -105.3522
11.7796 | -105.3587
11.7833 | -105.364
11.787 | -105.3722
11.7906 | -105.3793
11.7943 | -105.3806
11.798 | -105.3796
11.8017 | -105.3812
11.8053 | -105.381
11.809 | -105.3823
11.8127 | -105.3856
11.8163 | -105.3874
11.82 | -105.3889
11.8237 | -105.3893
11.8274 | -105.3895
11.831 | -105.3907
11.8347 | -105.3934
11.8384 | -105.3949
11.8421 | -105.3973
11.8457 | -105.4037
11.8494 | -105.41
11.8531 | -105.421
11.8567 | -105.4355
11.8604 | -105.4509
11.8641 | -105.4672
11.8678 | -105.4834
11.8714 | -105.4974
11.8751 | -105.5172
11.8788 | -105.5329
11.8825 | -105.5483
11.8861 | -105.565
11.8898 | -105.5813
11.8935 | -105.5962
11.8972 | -105.6078
11.9008 | -105.6183
11.9045 | -105.6231
11.9082 | -105.6234
11.9118 | -105.622
11.9155 | -105.6214
11.9192 | -105.6158
11.9229 | -105.6093
11.9265 | -105.6
11.9302 | -105.5916
11.9339 | -105.5853
11.9376 | -105.5753
11.9412 | -105.5679
11.9449 | -105.5636
11.9486 | -105.5592
11.9522 | -105.5565
11.9559 | -105.5551
11.9596 | -105.5531
11.9633 | -105.5497
11.9669 | -105.5467
11.9706 | -105.5467
11.9743 | -105.5461
11.978 | -105.5434
11.9816 | -105.5412
11.9853 | -105.5417
11.989 | -105.5416
11.9927 | -105.5395
11.9963 | -105.5339
12 | -105.524
12.0037 | -105.5136
12.0073 | -105.5062
12.011 | -105.5045
12.0147 | -105.501
12.0184 | -105.4935
12.022 | -105.4885
12.0257 | -105.4818
12.0294 | -105.4729
12.0331 | -105.4664
12.0367 | -105.4603
12.0404 | -105.4545
12.0441 | -105.4505
12.0478 | -105.447
12.0514 | -105.4463
12.0551 | -105.4476
12.0588 | -105.4497
12.0624 | -105.4503
12.0661 | -105.4506
12.0698 | -105.4573
12.0735 | -105.4681
12.0771 | -105.4823
12.0808 | -105.4961
12.0845 | -105.5083
12.0882 | -105.5227
12.0918 | -105.5446
12.0955 | -105.5664
12.0992 | -105.5889
12.1028 | -105.6071
12.1065 | -105.6235
12.1102 | -105.6418
12.1139 | -105.6551
12.1175 | -105.667
12.1212 | -105.6815
12.1249 | -105.6968
12.1286 | -105.7101
12.1322 | -105.7247
12.1359 | -105.739
12.1396 | -105.7499
12.1433 | -105.7572
12.1469 | -105.7611
12.1506 | -105.7667
12.1543 | -105.7704
12.1579 | -105.7719
12.1616 | -105.7719
12.1653 | -105.7755
12.169 | -105.7796
12.1726 | -105.7822
12.1763 | -105.7805
12.18 | -105.7748
12.1837 | -105.7637
12.1873 | -105.7558
12.191 | -105.7468
12.1947 | -105.741
12.1983 | -105.7323
12.202 | -105.7212
12.2057 | -105.7133
12.2094 | -105.7033
12.213 | -105.6926
12.2167 | -105.6833
12.2204 | -105.675
12.2241 | -105.6613
12.2277 | -105.6461
12.2314 | -105.6324
12.2351 | -105.6186
12.2388 | -105.6045
12.2424 | -105.5921
12.2461 | -105.5778
12.2498 | -105.5613
12.2534 | -105.5466
12.2571 | -105.5361
12.2608 | -105.5272
12.2645 | -105.5173
12.2681 | -105.505
12.2718 | -105.4956
12.2755 | -105.4938
12.2792 | -105.4965
12.2828 | -105.4985
12.2865 | -105.4992
12.2902 | -105.5048
12.2938 | -105.5139
12.2975 | -105.5202
12.3012 | -105.5263
12.3049 | -105.5334
12.3085 | -105.5404
12.3122 | -105.5483
12.3159 | -105.559
12.3196 | -105.5734
12.3232 | -105.5872
12.3269 | -105.5966
12.3306 | -105.6059
12.3343 | -105.6162
12.3379 | -105.6265
12.3416 | -105.6348
12.3453 | -105.6436
12.3489 | -105.6528
12.3526 | -105.6632
12.3563 | -105.6714
12.36 | -105.6782
12.3636 | -105.6851
12.3673 | -105.6921
12.371 | -105.7001
12.3747 | -105.7078
12.3783 | -105.7172
12.382 | -105.7261
12.3857 | -105.7338
12.3893 | -105.7411
12.393 | -105.7502
12.3967 | -105.7581
12.4004 | -105.7632
12.404 | -105.7717
12.4077 | -105.7809
12.4114 | -105.7884
12.4151 | -105.7978
12.4187 | -105.8064
12.4224 | -105.8147
12.4261 | -105.8218
12.4298 | -105.8289
12.4334 | -105.8349
12.4371 | -105.8387
12.4408 | -105.8437
12.4444 | -105.848
12.4481 | -105.854
12.4518 | -105.8579
12.4555 | -105.8513
12.4591 | -105.8431
12.4628 | -105.8368
12.4665 | -105.8313
12.4702 | -105.8197
12.4738 | -105.805
12.4775 | -105.7906
12.4812 | -105.7772
12.4848 | -105.7609
12.4885 | -105.7415
12.4922 | -105.7212
12.4959 | -105.7017
12.4995 | -105.6798
12.5032 | -105.6593
12.5069 | -105.6384
12.5106 | -105.6188
12.5142 | -105.6006
12.5179 | -105.5867
12.5216 | -105.5732
12.5253 | -105.5622
12.5289 | -105.5551
12.5326 | -105.5486
12.5363 | -105.5449
12.5399 | -105.548
12.5436 | -105.5525
12.5473 | -105.56
12.551 | -105.5697
12.5546 | -105.583
12.5583 | -105.599
12.562 | -105.6184
12.5657 | -105.641
12.5693 | -105.6633
12.573 | -105.6832
12.5767 | -105.7047
12.5803 | -105.7305
12.584 | -105.7547
12.5877 | -105.7764
12.5914 | -105.7976
12.595 | -105.816
12.5987 | -105.8347
12.6024 | -105.8522
12.6061 | -105.869
12.6097 | -105.8803
12.6134 | -105.8872
12.6171 | -105.8939
12.6208 | -105.8963
12.6244 | -105.8968
12.6281 | -105.8947
12.6318 | -105.8918
12.6354 | -105.8883
12.6391 | -105.8807
12.6428 | -105.871
12.6465 | -105.8546
12.6501 | -105.8393
12.6538 | -105.8243
12.6575 | -105.8096
12.6612 | -105.7911
12.6648 | -105.7766
12.6685 | -105.7646
12.6722 | -105.7499
12.6758 | -105.7343
12.6795 | -105.7177
12.6832 | -105.7035
12.6869 | -105.6919
12.6905 | -105.6819
12.6942 | -105.6674
12.6979 | -105.653
12.7016 | -105.6422
12.7052 | -105.6335
12.7089 | -105.6223
12.7126 | -105.6124
12.7163 | -105.6086
12.7199 | -105.6022
12.7236 | -105.5956
12.7273 | -105.5916
12.7309 | -105.5888
12.7346 | -105.5894
12.7383 | -105.5902
12.742 | -105.5925
12.7456 | -105.5997
12.7493 | -105.6068
12.753 | -105.6178
12.7567 | -105.6331
12.7603 | -105.6474
12.764 | -105.6601
12.7677 | -105.677
12.7713 | -105.6971
12.775 | -105.7148
12.7787 | -105.7346
12.7824 | -105.7552
12.786 | -105.7757
12.7897 | -105.7947
12.7934 | -105.8163
12.7971 | -105.8348
12.8007 | -105.8526
12.8044 | -105.8715
12.8081 | -105.8868
12.8118 | -105.9022
12.8154 | -105.9205
12.8191 | -105.9372
12.8228 | -105.9481
12.8264 | -105.9593
12.8301 | -105.9698
12.8338 | -105.977
12.8375 | -105.98
12.8411 | -105.9791
12.8448 | -105.9798
12.8485 | -105.9775
12.8522 | -105.9721
12.8558 | -105.9687
12.8595 | -105.9631
12.8632 | -105.9565
12.8668 | -105.9447
12.8705 | -105.9321
12.8742 | -105.9145
12.8779 | -105.901
12.8815 | -105.8848
12.8852 | -105.8659
12.8889 | -105.8468
12.8926 | -105.8318
12.8962 | -105.8162
12.8999 | -105.8035
12.9036 | -105.7905
12.9073 | -105.7792
12.9109 | -105.7694
12.9146 | -105.7581
12.9183 | -105.7503
12.9219 | -105.7455
12.9256 | -105.7441
12.9293 | -105.7429
12.933 | -105.7408
12.9366 | -105.7422
12.9403 | -105.745
12.944 | -105.7464
12.9477 | -105.7509
12.9513 | -105.7577
12.955 | -105.762
12.9587 | -105.7663
12.9624 | -105.7713
12.966 | -105.7776
12.9697 | -105.7837
12.9734 | -105.787
12.977 | -105.7909
12.9807 | -105.795
12.9844 | -105.8025
12.9881 | -105.8058
12.9917 | -105.8057
12.9954 | -105.8066
12.9991 | -105.8064
13.0028 | -105.8053
13.0064 | -105.8069
13.0101 | -105.8119
13.0138 | -105.8111
13.0174 | -105.8113
13.0211 | -105.8114
13.0248 | -105.8122
13.0285 | -105.8159
13.0321 | -105.8167
13.0358 | -105.8188
13.0395 | -105.8246
13.0432 | -105.8317
13.0468 | -105.8415
13.0505 | -105.8542
13.0542 | -105.8696
13.0579 | -105.8858
13.0615 | -105.9078
13.0652 | -105.9302
13.0689 | -105.9525
13.0725 | -105.9756
13.0762 | -105.9977
13.0799 | -106.0218
13.0836 | -106.0462
13.0872 | -106.069
13.0909 | -106.094
13.0946 | -106.1187
13.0983 | -106.1391
13.1019 | -106.1605
13.1056 | -106.1796
13.1093 | -106.1937
13.1129 | -106.2051
13.1166 | -106.2159
13.1203 | -106.2241
13.124 | -106.2269
13.1276 | -106.2283
13.1313 | -106.2296
13.135 | -106.2266
13.1387 | -106.2181
13.1423 | -106.208
13.146 | -106.1987
13.1497 | -106.1875
13.1534 | -106.1715
13.157 | -106.1547
13.1607 | -106.1408
13.1644 | -106.1302
13.168 | -106.1178
13.1717 | -106.0998
13.1754 | -106.0823
13.1791 | -106.0673
13.1827 | -106.055
13.1864 | -106.0442
13.1901 | -106.0312
13.1938 | -106.0192
13.1974 | -106.0053
13.2011 | -105.9958
13.2048 | -105.9877
13.2084 | -105.9779
13.2121 | -105.9682
13.2158 | -105.9559
13.2195 | -105.946
13.2231 | -105.9364
13.2268 | -105.9276
13.2305 | -105.9168
13.2342 | -105.9067
13.2378 | -105.9015
13.2415 | -105.8967
13.2452 | -105.8959
13.2489 | -105.8941
13.2525 | -105.8907
13.2562 | -105.8884
13.2599 | -105.8904
13.2635 | -105.8922
13.2672 | -105.8946
13.2709 | -105.8945
13.2746 | -105.8962
13.2782 | -105.8968
13.2819 | -105.8968
13.2856 | -105.9013
13.2893 | -105.907
13.2929 | -105.9129
13.2966 | -105.9159
13.3003 | -105.9198
13.3039 | -105.9262
13.3076 | -105.9322
13.3113 | -105.9376
13.315 | -105.9409
13.3186 | -105.9439
13.3223 | -105.9479
13.326 | -105.9514
13.3297 | -105.9557
13.3333 | -105.9592
13.337 | -105.958
13.3407 | -105.9578
13.3444 | -105.9563
13.348 | -105.9549
13.3517 | -105.9513
13.3554 | -105.9427
13.359 | -105.9336
13.3627 | -105.9279
13.3664 | -105.9227
13.3701 | -105.9179
13.3737 | -105.9115
13.3774 | -105.9028
13.3811 | -105.8974
13.3848 | -105.8915
13.3884 | -105.8866
13.3921 | -105.8833
13.3958 | -105.8802
13.3994 | -105.8798
13.4031 | -105.882
13.4068 | -105.884
13.4105 | -105.8856
13.4141 | -105.8851
13.4178 | -105.8843
13.4215 | -105.8852
13.4252 | -105.8858
13.4288 | -105.8831
13.4325 | -105.8812
13.4362 | -105.8787
13.4399 | -105.8755
13.4435 | -105.8721
13.4472 | -105.8703
13.4509 | -105.869
13.4545 | -105.8666
13.4582 | -105.8655
13.4619 | -105.8657
13.4656 | -105.8662
13.4692 | -105.8675
13.4729 | -105.8705
13.4766 | -105.8747
13.4803 | -105.8784
13.4839 | -105.8848
13.4876 | -105.892
13.4913 | -105.9038
13.4949 | -105.9165
13.4986 | -105.9282
13.5023 | -105.9407
13.506 | -105.9524
13.5096 | -105.9666
13.5133 | -105.981
13.517 | -105.9946
13.5207 | -106.0077
13.5243 | -106.0217
13.528 | -106.034
13.5317 | -106.0482
13.5354 | -106.0606
13.539 | -106.0722
13.5427 | -106.0787
13.5464 | -106.0853
13.55 | -106.0962
13.5537 | -106.1054
13.5574 | -106.1132
13.5611 | -106.1158
13.5647 | -106.1143
13.5684 | -106.114
13.5721 | -106.1129
13.5758 | -106.109
13.5794 | -106.1086
13.5831 | -106.1086
13.5868 | -106.1058
13.5905 | -106.1031
13.5941 | -106.0985
13.5978 | -106.0939
13.6015 | -106.09
13.6051 | -106.0855
13.6088 | -106.0809
13.6125 | -106.0762
13.6162 | -106.0718
13.6198 | -106.0639
13.6235 | -106.0568
13.6272 | -106.0528
13.6309 | -106.0482
13.6345 | -106.0437
13.6382 | -106.039
13.6419 | -106.0319
13.6455 | -106.024
13.6492 | -106.0169
13.6529 | -106.0128
13.6566 | -106.0065
13.6602 | -105.9991
13.6639 | -105.99
13.6676 | -105.9811
13.6713 | -105.9726
13.6749 | -105.9628
13.6786 | -105.9542
13.6823 | -105.9449
13.686 | -105.9352
13.6896 | -105.9274
13.6933 | -105.9193
13.697 | -105.9097
13.7006 | -105.8991
13.7043 | -105.8907
13.708 | -105.8812
13.7117 | -105.872
13.7153 | -105.8643
13.719 | -105.8553
13.7227 | -105.8454
13.7264 | -105.8358
13.73 | -105.8279
13.7337 | -105.8219
13.7374 | -105.819
13.741 | -105.8164
13.7447 | -105.8092
13.7484 | -105.8084
13.7521 | -105.8087
13.7557 | -105.8115
13.7594 | -105.8141
13.7631 | -105.8179
13.7668 | -105.8228
13.7704 | -105.8282
13.7741 | -105.8381
13.7778 | -105.8438
13.7815 | -105.8506
13.7851 | -105.8577
13.7888 | -105.8659
13.7925 | -105.8747
13.7961 | -105.8824
13.7998 | -105.8929
13.8035 | -105.9041
13.8072 | -105.9155
13.8108 | -105.9243
13.8145 | -105.9321
13.8182 | -105.9417
13.8219 | -105.9522
13.8255 | -105.9594
13.8292 | -105.9654
13.8329 | -105.972
13.8365 | -105.9772
13.8402 | -105.9806
13.8439 | -105.987
13.8476 | -105.9926
13.8512 | -105.9921
13.8549 | -105.9922
13.8586 | -105.9909
13.8623 | -105.9889
13.8659 | -105.9852
13.8696 | -105.978
13.8733 | -105.9705
13.877 | -105.9615
13.8806 | -105.9523
13.8843 | -105.942
13.888 | -105.9308
13.8916 | -105.9195
13.8953 | -105.9086
13.899 | -105.8996
13.9027 | -105.8873
13.9063 | -105.8754
13.91 | -105.8625
13.9137 | -105.8496
13.9174 | -105.8406
13.921 | -105.8354
13.9247 | -105.83
13.9284 | -105.8253
13.932 | -105.8211
13.9357 | -105.8204
13.9394 | -105.8221
13.9431 | -105.8274
13.9467 | -105.83
13.9504 | -105.8354
13.9541 | -105.8424
13.9578 | -105.8491
13.9614 | -105.8459
13.9651 | -105.8437
13.9688 | -105.85
13.9725 | -105.8562
13.9761 | -105.8605
13.9798 | -105.8635
13.9835 | -105.8656
13.9871 | -105.8656
13.9908 | -105.8636
13.9945 | -105.8602
13.9982 | -105.8538
14.0018 | -105.8434
14.0055 | -105.8297
14.0092 | -105.816
14.0129 | -105.7984
14.0165 | -105.778
14.0202 | -105.7553
14.0239 | -105.7305
14.0275 | -105.7048
14.0312 | -105.6767
14.0349 | -105.6493
14.0386 | -105.6249
14.0422 | -105.6028
14.0459 | -105.5787
14.0496 | -105.5572
14.0533 | -105.5399
14.0569 | -105.5253
14.0606 | -105.5141
14.0643 | -105.5037
14.068 | -105.499
14.0716 | -105.5001
14.0753 | -105.5042
14.079 | -105.5126
14.0826 | -105.524
14.0863 | -105.5391
14.09 | -105.5564
14.0937 | -105.5761
14.0973 | -105.6002
14.101 | -105.627
14.1047 | -105.6573
14.1084 | -105.6897
14.112 | -105.7225
14.1157 | -105.7562
14.1194 | -105.7872
14.123 | -105.8192
14.1267 | -105.8508
14.1304 | -105.8821
14.1341 | -105.9129
14.1377 | -105.9443
14.1414 | -105.9756
14.1451 | -106.0052
14.1488 | -106.0333
14.1524 | -106.0585
14.1561 | -106.0839
14.1598 | -106.1118
14.1635 | -106.1456
14.1671 | -106.1739
14.1708 | -106.1921
14.1745 | -106.2071
14.1781 | -106.2216
14.1818 | -106.2355
14.1855 | -106.2469
14.1892 | -106.2569
14.1928 | -106.2649
14.1965 | -106.2722
14.2002 | -106.2798
14.2039 | -106.2879
14.2075 | -106.2958
14.2112 | -106.2995
14.2149 | -106.3051
14.2185 | -106.3121
14.2222 | -106.3164
14.2259 | -106.3182
14.2296 | -106.3209
14.2332 | -106.3269
14.2369 | -106.331
14.2406 | -106.3296
14.2443 | -106.3278
14.2479 | -106.3272
14.2516 | -106.3264
14.2553 | -106.3263
14.259 | -106.3269
14.2626 | -106.3257
14.2663 | -106.3246
14.27 | -106.3215
14.2736 | -106.3142
14.2773 | -106.3094
14.281 | -106.3063
14.2847 | -106.3037
14.2883 | -106.2999
14.292 | -106.2968
14.2957 | -106.293
14.2994 | -106.2824
14.303 | -106.2685
14.3067 | -106.2511
14.3104 | -106.2357
14.314 | -106.2211
14.3177 | -106.209
14.3214 | -106.1951
14.3251 | -106.1825
14.3287 | -106.1682
14.3324 | -106.1529
14.3361 | -106.1371
14.3398 | -106.1232
14.3434 | -106.1041
14.3471 | -106.0843
14.3508 | -106.0684
14.3545 | -106.051
14.3581 | -106.0358
14.3618 | -106.0208
14.3655 | -106.0021
14.3691 | -105.9857
14.3728 | -105.9701
14.3765 | -105.9552
14.3802 | -105.9418
14.3838 | -105.9285
14.3875 | -105.9156
14.3912 | -105.9049
14.3949 | -105.8968
14.3985 | -105.8891
14.4022 | -105.8827
14.4059 | -105.8773
14.4095 | -105.8721
14.4132 | -105.867
14.4169 | -105.8634
14.4206 | -105.8598
14.4242 | -105.8561
14.4279 | -105.855
14.4316 | -105.853
14.4353 | -105.8521
14.4389 | -105.8494
14.4426 | -105.848
14.4463 | -105.8505
14.45 | -105.8515
14.4536 | -105.8501
14.4573 | -105.8488
14.461 | -105.8481
14.4646 | -105.8462
14.4683 | -105.8479
14.472 | -105.8489
14.4757 | -105.85
14.4793 | -105.8529
14.483 | -105.8516
14.4867 | -105.8499
14.4904 | -105.8515
14.494 | -105.8569
14.4977 | -105.8622
14.5014 | -105.87
14.5051 | -105.8814
14.5087 | -105.8918
14.5124 | -105.9009
14.5161 | -105.909
14.5197 | -105.9207
14.5234 | -105.9315
14.5271 | -105.9428
14.5308 | -105.9574
14.5344 | -105.9727
14.5381 | -105.9857
14.5418 | -106.0017
14.5455 | -106.0195
14.5491 | -106.0358
14.5528 | -106.0548
14.5565 | -106.0738
14.5601 | -106.0907
14.5638 | -106.1112
14.5675 | -106.1292
14.5712 | -106.1461
14.5748 | -106.1634
14.5785 | -106.1771
14.5822 | -106.1891
14.5859 | -106.2007
14.5895 | -106.2137
14.5932 | -106.2251
14.5969 | -106.2373
14.6006 | -106.2488
14.6042 | -106.2575
14.6079 | -106.2649
14.6116 | -106.2705
14.6152 | -106.2754
14.6189 | -106.2813
14.6226 | -106.2875
14.6263 | -106.2947
14.6299 | -106.303
14.6336 | -106.3094
14.6373 | -106.3141
14.641 | -106.3217
14.6446 | -106.3298
14.6483 | -106.3378
14.652 | -106.3443
14.6556 | -106.3505
14.6593 | -106.3579
14.663 | -106.3672
14.6667 | -106.3757
14.6703 | -106.3804
14.674 | -106.386
14.6777 | -106.3907
14.6814 | -106.3968
14.685 | -106.4026
14.6887 | -106.4039
14.6924 | -106.406
14.6961 | -106.4084
14.6997 | -106.409
14.7034 | -106.4067
14.7071 | -106.4056
14.7107 | -106.4028
14.7144 | -106.3989
14.7181 | -106.3935
14.7218 | -106.3875
14.7254 | -106.3806
14.7291 | -106.3722
14.7328 | -106.3628
14.7365 | -106.3521
14.7401 | -106.34
14.7438 | -106.3233
14.7475 | -106.3099
14.7511 | -106.2963
14.7548 | -106.2823
14.7585 | -106.2685
14.7622 | -106.2541
14.7658 | -106.2403
14.7695 | -106.225
14.7732 | -106.2107
14.7769 | -106.199
14.7805 | -106.1894
14.7842 | -106.1818
14.7879 | -106.1757
14.7916 | -106.1726
14.7952 | -106.1692
14.7989 | -106.1661
14.8026 | -106.166
14.8062 | -106.165
14.8099 | -106.1646
14.8136 | -106.1679
14.8173 | -106.1763
14.8209 | -106.1855
14.8246 | -106.1954
14.8283 | -106.2061
14.832 | -106.2193
14.8356 | -106.2336
14.8393 | -106.2431
14.843 | -106.2518
14.8466 | -106.2599
14.8503 | -106.2689
14.854 | -106.2788
14.8577 | -106.2863
14.8613 | -106.2927
14.865 | -106.2984
14.8687 | -106.3027
14.8724 | -106.3086
14.876 | -106.3122
14.8797 | -106.3129
14.8834 | -106.3138
14.8871 | -106.3125
14.8907 | -106.3077
14.8944 | -106.304
14.8981 | -106.3007
14.9017 | -106.3005
14.9054 | -106.2997
14.9091 | -106.2956
14.9128 | -106.2907
14.9164 | -106.2861
14.9201 | -106.283
14.9238 | -106.2808
14.9275 | -106.2791
14.9311 | -106.278
14.9348 | -106.275
14.9385 | -106.2713
14.9421 | -106.2699
14.9458 | -106.2657
14.9495 | -106.2633
14.9532 | -106.2606
14.9568 | -106.2609
14.9605 | -106.2611
14.9642 | -106.2619
14.9679 | -106.2611
14.9715 | -106.2583
14.9752 | -106.2579
14.9789 | -106.2593
14.9826 | -106.256
14.9862 | -106.2509
14.9899 | -106.2486
14.9936 | -106.2446
14.9972 | -106.2389
15.0009 | -106.235
15.0046 | -106.2317
15.0083 | -106.2272
15.0119 | -106.2211
15.0156 | -106.2143
15.0193 | -106.2071
15.023 | -106.2006
15.0266 | -106.1934
15.0303 | -106.1836
15.034 | -106.1748
15.0376 | -106.1649
15.0413 | -106.1546
15.045 | -106.145
15.0487 | -106.1349
15.0523 | -106.1275
15.056 | -106.1229
15.0597 | -106.1162
15.0634 | -106.1142
15.067 | -106.1116
15.0707 | -106.1065
15.0744 | -106.1047
15.0781 | -106.1084
15.0817 | -106.1102
15.0854 | -106.114
15.0891 | -106.1183
15.0927 | -106.1252
15.0964 | -106.1357
15.1001 | -106.1475
15.1038 | -106.1589
15.1074 | -106.1726
15.1111 | -106.185
15.1148 | -106.1967
15.1185 | -106.2072
15.1221 | -106.2187
15.1258 | -106.2312
15.1295 | -106.2424
15.1332 | -106.2519
15.1368 | -106.2607
15.1405 | -106.2659
15.1442 | -106.272
15.1478 | -106.2752
15.1515 | -106.2756
15.1552 | -106.2758
15.1589 | -106.271
15.1625 | -106.2689
15.1662 | -106.2648
15.1699 | -106.2591
15.1736 | -106.2504
15.1772 | -106.24
15.1809 | -106.2278
15.1846 | -106.2153
15.1882 | -106.2052
15.1919 | -106.1946
15.1956 | -106.183
15.1993 | -106.1755
15.2029 | -106.1654
15.2066 | -106.1515
15.2103 | -106.1386
15.214 | -106.1272
15.2176 | -106.1152
15.2213 | -106.1039
15.225 | -106.0952
15.2287 | -106.0883
15.2323 | -106.0811
15.236 | -106.0763
15.2397 | -106.0725
15.2433 | -106.0695
15.247 | -106.0669
15.2507 | -106.0636
15.2544 | -106.0645
15.258 | -106.066
15.2617 | -106.0674
15.2654 | -106.0694
15.2691 | -106.0726
15.2727 | -106.0757
15.2764 | -106.0805
15.2801 | -106.084
15.2837 | -106.0903
15.2874 | -106.0978
15.2911 | -106.1041
15.2948 | -106.1121
15.2984 | -106.122
15.3021 | -106.132
15.3058 | -106.1406
15.3095 | -106.148
15.3131 | -106.1549
15.3168 | -106.1642
15.3205 | -106.1741
15.3242 | -106.1833
15.3278 | -106.1915
15.3315 | -106.2008
15.3352 | -106.2116
15.3388 | -106.2201
15.3425 | -106.2267
15.3462 | -106.2324
15.3499 | -106.2344
15.3535 | -106.2365
15.3572 | -106.2413
15.3609 | -106.2429
15.3646 | -106.2372
15.3682 | -106.2395
15.3719 | -106.2421
15.3756 | -106.2414
15.3792 | -106.2378
15.3829 | -106.2314
15.3866 | -106.2241
15.3903 | -106.217
15.3939 | -106.2099
15.3976 | -106.1999
15.4013 | -106.1902
15.405 | -106.1818
15.4086 | -106.1744
15.4123 | -106.1664
15.416 | -106.1578
15.4197 | -106.1509
15.4233 | -106.1442
15.427 | -106.1368
15.4307 | -106.1298
15.4343 | -106.1238
15.438 | -106.1184
15.4417 | -106.1116
15.4454 | -106.1057
15.449 | -106.1001
15.4527 | -106.0964
15.4564 | -106.0937
15.4601 | -106.092
15.4637 | -106.0902
15.4674 | -106.0885
15.4711 | -106.0888
15.4747 | -106.0868
15.4784 | -106.0852
15.4821 | -106.0858
15.4858 | -106.0868
15.4894 | -106.0875
15.4931 | -106.0867
15.4968 | -106.0875
15.5005 | -106.0881
15.5041 | -106.0871
15.5078 | -106.0866
15.5115 | -106.0851
15.5152 | -106.0813
15.5188 | -106.0779
15.5225 | -106.0717
15.5262 | -106.0674
15.5298 | -106.0633
15.5335 | -106.0564
15.5372 | -106.0505
15.5409 | -106.0435
15.5445 | -106.0346
15.5482 | -106.0276
15.5519 | -106.0159
15.5556 | -106.0034
15.5592 | -105.9921
15.5629 | -105.9813
15.5666 | -105.9726
15.5702 | -105.9674
15.5739 | -105.9589
15.5776 | -105.9506
15.5813 | -105.9436
15.5849 | -105.9357
15.5886 | -105.9299
15.5923 | -105.924
15.596 | -105.919
15.5996 | -105.9166
15.6033 | -105.9173
15.607 | -105.9199
15.6107 | -105.9244
15.6143 | -105.9285
15.618 | -105.9348
15.6217 | -105.9429
15.6253 | -105.9525
15.629 | -105.9612
15.6327 | -105.9686
15.6364 | -105.9763
15.64 | -105.9848
15.6437 | -105.9932
15.6474 | -106.0056
15.6511 | -106.0174
15.6547 | -106.0305
15.6584 | -106.0446
15.6621 | -106.0582
15.6657 | -106.0736
15.6694 | -106.0891
15.6731 | -106.1031
15.6768 | -106.1152
15.6804 | -106.1287
15.6841 | -106.138
15.6878 | -106.1463
15.6915 | -106.1557
15.6951 | -106.1652
15.6988 | -106.1712
15.7025 | -106.176
15.7062 | -106.1757
15.7098 | -106.1762
15.7135 | -106.1742
15.7172 | -106.1728
15.7208 | -106.1683
15.7245 | -106.1622
15.7282 | -106.1559
15.7319 | -106.15
15.7355 | -106.1431
15.7392 | -106.1351
15.7429 | -106.126
15.7466 | -106.1171
15.7502 | -106.1085
15.7539 | -106.0996
15.7576 | -106.0921
15.7612 | -106.0822
15.7649 | -106.0747
15.7686 | -106.0684
15.7723 | -106.0608
15.7759 | -106.0551
15.7796 | -106.0499
15.7833 | -106.0434
15.787 | -106.0371
15.7906 | -106.0326
15.7943 | -106.0265
15.798 | -106.0235
15.8017 | -106.0193
15.8053 | -106.0141
15.809 | -106.0113
15.8127 | -106.0111
15.8163 | -106.0064
15.82 | -106.0012
15.8237 | -105.996
15.8274 | -105.9914
15.831 | -105.9863
15.8347 | -105.9816
15.8384 | -105.9791
15.8421 | -105.9759
15.8457 | -105.9716
15.8494 | -105.9655
15.8531 | -105.9598
15.8567 | -105.954
15.8604 | -105.9518
15.8641 | -105.9472
15.8678 | -105.9392
15.8714 | -105.9293
15.8751 | -105.9223
15.8788 | -105.9151
15.8825 | -105.905
15.8861 | -105.8952
15.8898 | -105.8881
15.8935 | -105.8811
15.8972 | -105.873
15.9008 | -105.8651
15.9045 | -105.8569
15.9082 | -105.8468
15.9118 | -105.8392
15.9155 | -105.8285
15.9192 | -105.8182
15.9229 | -105.8096
15.9265 | -105.8031
15.9302 | -105.7944
15.9339 | -105.7898
15.9376 | -105.7827
15.9412 | -105.776
15.9449 | -105.7682
15.9486 | -105.762
15.9522 | -105.7551
15.9559 | -105.749
15.9596 | -105.7425
15.9633 | -105.7385
15.9669 | -105.7329
15.9706 | -105.7263
15.9743 | -105.7214
15.978 | -105.7177
15.9816 | -105.715
15.9853 | -105.7119
15.989 | -105.7089
15.9927 | -105.7075
15.9963 | -105.7062
16 | -105.7056
16.0037 | -105.7062
16.0073 | -105.7074
16.011 | -105.7107
16.0147 | -105.7111
16.0184 | -105.7095
16.022 | -105.7107
16.0257 | -105.7104
16.0294 | -105.7111
16.0331 | -105.7135
16.0367 | -105.7172
16.0404 | -105.7202
16.0441 | -105.727
16.0478 | -105.7356
16.0514 | -105.742
16.0551 | -105.7466
16.0588 | -105.7532
16.0624 | -105.7578
16.0661 | -105.7611
16.0698 | -105.7629
16.0735 | -105.7695
16.0771 | -105.7789
16.0808 | -105.7876
16.0845 | -105.7946
16.0882 | -105.7995
16.0918 | -105.8069
16.0955 | -105.8166
16.0992 | -105.8267
16.1028 | -105.8354
16.1065 | -105.8422
16.1102 | -105.8508
16.1139 | -105.859
16.1175 | -105.8644
16.1212 | -105.8704
16.1249 | -105.881
16.1286 | -105.8917
16.1322 | -105.8998
16.1359 | -105.9108
16.1396 | -105.9212
16.1433 | -105.93
16.1469 | -105.9418
16.1506 | -105.9526
16.1543 | -105.9592
16.1579 | -105.9666
16.1616 | -105.9749
16.1653 | -105.9797
16.169 | -105.985
16.1726 | -105.9907
16.1763 | -105.9967
16.18 | -106.0025
16.1837 | -106.007
16.1873 | -106.0099
16.191 | -106.0133
16.1947 | -106.016
16.1983 | -106.0202
16.202 | -106.0252
16.2057 | -106.0289
16.2094 | -106.0277
16.213 | -106.0265
16.2167 | -106.0288
16.2204 | -106.0327
16.2241 | -106.0351
16.2277 | -106.037
16.2314 | -106.0382
16.2351 | -106.0428
16.2388 | -106.0481
16.2424 | -106.051
16.2461 | -106.0583
16.2498 | -106.0666
16.2534 | -106.0735
16.2571 | -106.0811
16.2608 | -106.0911
16.2645 | -106.1016
16.2681 | -106.1108
16.2718 | -106.117
16.2755 | -106.124
16.2792 | -106.132
16.2828 | -106.1436
16.2865 | -106.1582
16.2902 | -106.1711
16.2938 | -106.1854
16.2975 | -106.2004
16.3012 | -106.219
16.3049 | -106.2335
16.3085 | -106.2471
16.3122 | -106.2638
16.3159 | -106.2839
16.3196 | -106.3017
16.3232 | -106.3185
16.3269 | -106.3348
16.3306 | -106.3495
16.3343 | -106.3653
16.3379 | -106.3815
16.3416 | -106.3962
16.3453 | -106.4111
16.3489 | -106.4258
16.3526 | -106.4405
16.3563 | -106.4517
16.36 | -106.464
16.3636 | -106.4778
16.3673 | -106.4885
16.371 | -106.4977
16.3747 | -106.5083
16.3783 | -106.5186
16.382 | -106.5252
16.3857 | -106.5321
16.3893 | -106.5379
16.393 | -106.5436
16.3967 | -106.5498
16.4004 | -106.5546
16.404 | -106.5591
16.4077 | -106.5649
16.4114 | -106.5686
16.4151 | -106.5732
16.4187 | -106.5794
16.4224 | -106.5822
16.4261 | -106.5859
16.4298 | -106.5893
16.4334 | -106.5932
16.4371 | -106.5951
16.4408 | -106.5962
16.4444 | -106.5997
16.4481 | -106.604
16.4518 | -106.6035
16.4555 | -106.6012
16.4591 | -106.6015
16.4628 | -106.6003
16.4665 | -106.5986
16.4702 | -106.5931
16.4738 | -106.5869
16.4775 | -106.5814
16.4812 | -106.5751
16.4848 | -106.5689
16.4885 | -106.5627
16.4922 | -106.553
16.4959 | -106.5382
16.4995 | -106.5229
16.5032 | -106.5072
16.5069 | -106.4958
16.5106 | -106.482
16.5142 | -106.4667
16.5179 | -106.4517
16.5216 | -106.4353
16.5253 | -106.4194
16.5289 | -106.4041
16.5326 | -106.3904
16.5363 | -106.3754
16.5399 | -106.3596
16.5436 | -106.3466
16.5473 | -106.3338
16.551 | -106.3217
16.5546 | -106.3131
16.5583 | -106.3067
16.562 | -106.3022
16.5657 | -106.2987
16.5693 | -106.2959
16.573 | -106.2965
16.5767 | -106.2993
16.5803 | -106.3015
16.584 | -106.3037
16.5877 | -106.3093
16.5914 | -106.3191
16.595 | -106.33
16.5987 | -106.3427
16.6024 | -106.3541
16.6061 | -106.3664
16.6097 | -106.381
16.6134 | -106.3966
16.6171 | -106.416
16.6208 | -106.436
16.6244 | -106.4563
16.6281 | -106.4763
16.6318 | -106.4972
16.6354 | -106.5184
16.6391 | -106.5393
16.6428 | -106.5583
16.6465 | -106.5767
16.6501 | -106.5948
16.6538 | -106.6098
16.6575 | -106.6267
16.6612 | -106.6431
16.6648 | -106.6581
16.6685 | -106.6703
16.6722 | -106.6808
16.6758 | -106.6913
16.6795 | -106.6989
16.6832 | -106.703
16.6869 | -106.7043
16.6905 | -106.7045
16.6942 | -106.7042
16.6979 | -106.7004
16.7016 | -106.6983
16.7052 | -106.6948
16.7089 | -106.6914
16.7126 | -106.6851
16.7163 | -106.6756
16.7199 | -106.6635
16.7236 | -106.65
16.7273 | -106.6377
16.7309 | -106.6281
16.7346 | -106.6178
16.7383 | -106.6065
16.742 | -106.5964
16.7456 | -106.5832
16.7493 | -106.5723
16.753 | -106.5607
16.7567 | -106.5478
16.7603 | -106.5364
16.764 | -106.5254
16.7677 | -106.5134
16.7714 | -106.5019
16.775 | -106.4914
16.7787 | -106.4842
16.7824 | -106.475
16.786 | -106.4695
16.7897 | -106.4626
16.7934 | -106.4554
16.7971 | -106.4504
16.8007 | -106.4472
16.8044 | -106.4429
16.8081 | -106.4404
16.8118 | -106.4376
16.8154 | -106.4361
16.8191 | -106.4357
16.8228 | -106.4385
16.8264 | -106.4378
16.8301 | -106.4351
16.8338 | -106.4339
16.8375 | -106.4331
16.8411 | -106.4361
16.8448 | -106.4396
16.8485 | -106.4416
16.8522 | -106.4448
16.8558 | -106.4454
16.8595 | -106.4461
16.8632 | -106.4498
16.8668 | -106.4517
16.8705 | -106.4544
16.8742 | -106.4599
16.8779 | -106.466
16.8815 | -106.4736
16.8852 | -106.4829
16.8889 | -106.4924
16.8926 | -106.5023
16.8962 | -106.5124
16.8999 | -106.5213
16.9036 | -106.533
16.9073 | -106.5441
16.9109 | -106.5562
16.9146 | -106.5707
16.9183 | -106.5866
16.9219 | -106.604
16.9256 | -106.6215
16.9293 | -106.6401
16.933 | -106.6586
16.9366 | -106.6758
16.9403 | -106.6939
16.944 | -106.712
16.9477 | -106.7269
16.9513 | -106.7438
16.955 | -106.7588
16.9587 | -106.7773
16.9624 | -106.7957
16.966 | -106.8128
16.9697 | -106.8267
16.9734 | -106.839
16.977 | -106.8513
16.9807 | -106.863
16.9844 | -106.8714
16.9881 | -106.8791
16.9917 | -106.8838
16.9954 | -106.887
16.9991 | -106.8881
17.0028 | -106.8887
17.0064 | -106.8883
17.0101 | -106.8876
17.0138 | -106.8856
17.0174 | -106.884
17.0211 | -106.8798
17.0248 | -106.8729
17.0285 | -106.8623
17.0321 | -106.8519
17.0358 | -106.842
17.0395 | -106.8314
17.0432 | -106.8191
17.0468 | -106.807
17.0505 | -106.7938
17.0542 | -106.7815
17.0579 | -106.7696
17.0615 | -106.759
17.0652 | -106.7439
17.0689 | -106.7303
17.0725 | -106.7187
17.0762 | -106.7048
17.0799 | -106.691
17.0836 | -106.6763
17.0872 | -106.663
17.0909 | -106.6523
17.0946 | -106.6398
17.0983 | -106.6297
17.1019 | -106.6193
17.1056 | -106.6072
17.1093 | -106.6008
17.1129 | -106.594
17.1166 | -106.5855
17.1203 | -106.5769
17.124 | -106.5679
17.1276 | -106.5621
17.1313 | -106.5554
17.135 | -106.5519
17.1387 | -106.5481
17.1423 | -106.5447
17.146 | -106.5418
17.1497 | -106.5405
17.1534 | -106.5382
17.157 | -106.5354
17.1607 | -106.5343
17.1644 | -106.5311
17.168 | -106.5296
17.1717 | -106.5293
17.1754 | -106.5271
17.1791 | -106.5267
17.1827 | -106.5267
17.1864 | -106.5285
17.1901 | -106.533
17.1938 | -106.5359
17.1974 | -106.5378
17.2011 | -106.5406
17.2048 | -106.5438
17.2084 | -106.5478
17.2121 | -106.5529
17.2158 | -106.5586
17.2195 | -106.567
17.2231 | -106.5757
17.2268 | -106.5846
17.2305 | -106.5929
17.2342 | -106.6019
17.2378 | -106.613
17.2415 | -106.6261
17.2452 | -106.6397
17.2489 | -106.649
17.2525 | -106.6595
17.2562 | -106.6689
17.2599 | -106.677
17.2635 | -106.6851
17.2672 | -106.6943
17.2709 | -106.7035
17.2746 | -106.7131
17.2782 | -106.7227
17.2819 | -106.7289
17.2856 | -106.7333
17.2893 | -106.7368
17.2929 | -106.7387
17.2966 | -106.7398
17.3003 | -106.7386
17.3039 | -106.7363
17.3076 | -106.7317
17.3113 | -106.7288
17.315 | -106.7247
17.3186 | -106.7169
17.3223 | -106.7118
17.326 | -106.7036
17.3297 | -106.6947
17.3333 | -106.6838
17.337 | -106.6692
17.3407 | -106.6583
17.3444 | -106.6441
17.348 | -106.6299
17.3517 | -106.6141
17.3554 | -106.5945
17.359 | -106.5746
17.3627 | -106.5558
17.3664 | -106.5401
17.3701 | -106.5269
17.3737 | -106.5101
17.3774 | -106.4949
17.3811 | -106.4789
17.3848 | -106.4631
17.3884 | -106.4495
17.3921 | -106.4353
17.3958 | -106.4203
17.3994 | -106.4088
17.4031 | -106.3974
17.4068 | -106.3843
17.4105 | -106.3749
17.4141 | -106.3654
17.4178 | -106.356
17.4215 | -106.3503
17.4252 | -106.3453
17.4288 | -106.3402
17.4325 | -106.3353
17.4362 | -106.3315
17.4399 | -106.331
17.4435 | -106.3318
17.4472 | -106.3326
17.4509 | -106.3345
17.4545 | -106.3399
17.4582 | -106.3473
17.4619 | -106.3562
17.4656 | -106.3622
17.4692 | -106.3698
17.4729 | -106.3796
17.4766 | -106.3918
17.4803 | -106.402
17.4839 | -106.4134
17.4876 | -106.4246
17.4913 | -106.4354
17.4949 | -106.4462
17.4986 | -106.4601
17.5023 | -106.4713
17.506 | -106.4838
17.5096 | -106.4941
17.5133 | -106.5038
17.517 | -106.5114
17.5207 | -106.5183
17.5243 | -106.5249
17.528 | -106.5296
17.5317 | -106.5355
17.5354 | -106.5399
17.539 | -106.5419
17.5427 | -106.5447
17.5464 | -106.5477
17.55 | -106.5502
17.5537 | -106.5508
17.5574 | -106.5484
17.5611 | -106.547
17.5647 | -106.545
17.5684 | -106.543
17.5721 | -106.54
17.5758 | -106.5387
17.5794 | -106.5377
17.5831 | -106.537
17.5868 | -106.5375
17.5905 | -106.5341
17.5941 | -106.5309
17.5978 | -106.5286
17.6015 | -106.5273
17.6051 | -106.5286
17.6088 | -106.5313
17.6125 | -106.5303
17.6162 | -106.5247
17.6198 | -106.5226
17.6235 | -106.5233
17.6272 | -106.5221
17.6309 | -106.5203
17.6345 | -106.5208
17.6382 | -106.5206
17.6419 | -106.5202
17.6455 | -106.5205
17.6492 | -106.5204
17.6529 | -106.52
17.6566 | -106.5213
17.6602 | -106.5199
17.6639 | -106.5203
17.6676 | -106.521
17.6713 | -106.5206
17.6749 | -106.5202
17.6786 | -106.5208
17.6823 | -106.5214
17.6859 | -106.5207
17.6896 | -106.5197
17.6933 | -106.5187
17.697 | -106.5175
17.7006 | -106.516
17.7043 | -106.5155
17.708 | -106.5149
17.7117 | -106.5172
17.7153 | -106.5222
17.719 | -106.5273
17.7227 | -106.5301
17.7264 | -106.5331
17.73 | -106.5377
17.7337 | -106.5419
17.7374 | -106.5472
17.741 | -106.5523
17.7447 | -106.5582
17.7484 | -106.5646
17.7521 | -106.5703
17.7557 | -106.5784
17.7594 | -106.5853
17.7631 | -106.5951
17.7668 | -106.6034
17.7704 | -106.6128
17.7741 | -106.6232
17.7778 | -106.6339
17.7815 | -106.6439
17.7851 | -106.6527
17.7888 | -106.6615
17.7925 | -106.6718
17.7961 | -106.6798
17.7998 | -106.6856
17.8035 | -106.692
17.8072 | -106.6992
17.8108 | -106.7076
17.8145 | -106.7131
17.8182 | -106.7177
17.8219 | -106.7216
17.8255 | -106.7233
17.8292 | -106.7259
17.8329 | -106.727
17.8365 | -106.7281
17.8402 | -106.7286
17.8439 | -106.7301
17.8476 | -106.7338
17.8512 | -106.7354
17.8549 | -106.7352
17.8586 | -106.7322
17.8623 | -106.7278
17.8659 | -106.7286
17.8696 | -106.7296
17.8733 | -106.7268
17.877 | -106.7222
17.8806 | -106.72
17.8843 | -106.718
17.888 | -106.7137
17.8916 | -106.7088
17.8953 | -106.7058
17.899 | -106.7036
17.9027 | -106.6992
17.9063 | -106.6943
17.91 | -106.6877
17.9137 | -106.681
17.9174 | -106.6726
17.921 | -106.6638
17.9247 | -106.6539
17.9284 | -106.6459
17.932 | -106.6366
17.9357 | -106.6263
17.9394 | -106.6154
17.9431 | -106.6049
17.9467 | -106.5958
17.9504 | -106.5856
17.9541 | -106.5759
17.9578 | -106.565
17.9614 | -106.5537
17.9651 | -106.5393
17.9688 | -106.5251
17.9725 | -106.5127
17.9761 | -106.5028
17.9798 | -106.4935
17.9835 | -106.486
17.9871 | -106.4781
17.9908 | -106.4693
17.9945 | -106.4608
17.9982 | -106.4546
18.0018 | -106.4494
18.0055 | -106.4424
18.0092 | -106.4365
18.0129 | -106.4326
18.0165 | -106.4368
18.0202 | -106.4339
18.0239 | -106.4323
18.0275 | -106.4296
18.0312 | -106.4281
18.0349 | -106.4291
18.0386 | -106.4327
18.0422 | -106.4365
18.0459 | -106.4417
18.0496 | -106.4476
18.0533 | -106.4543
18.0569 | -106.459
18.0606 | -106.4637
18.0643 | -106.4671
18.068 | -106.4729
18.0716 | -106.482
18.0753 | -106.4888
18.079 | -106.4957
18.0826 | -106.5023
18.0863 | -106.5072
18.09 | -106.5129
18.0937 | -106.5164
18.0973 | -106.5184
18.101 | -106.5205
18.1047 | -106.5226
18.1084 | -106.5231
18.112 | -106.5244
18.1157 | -106.5241
18.1194 | -106.5203
18.123 | -106.5138
18.1267 | -106.5086
18.1304 | -106.5026
18.1341 | -106.4957
18.1377 | -106.4879
18.1414 | -106.4798
18.1451 | -106.4707
18.1488 | -106.46
18.1524 | -106.4486
18.1561 | -106.4392
18.1598 | -106.4276
18.1635 | -106.4159
18.1671 | -106.4051
18.1708 | -106.3939
18.1745 | -106.3847
18.1781 | -106.3755
18.1818 | -106.3678
18.1855 | -106.3601
18.1892 | -106.3544
18.1928 | -106.3508
18.1965 | -106.348
18.2002 | -106.3457
18.2039 | -106.3445
18.2075 | -106.3457
18.2112 | -106.3478
18.2149 | -106.349
18.2185 | -106.3513
18.2222 | -106.3581
18.2259 | -106.3675
18.2296 | -106.3767
18.2332 | -106.3864
18.2369 | -106.3969
18.2406 | -106.4081
18.2443 | -106.419
18.2479 | -106.4306
18.2516 | -106.4417
18.2553 | -106.4516
18.259 | -106.4631
18.2626 | -106.4775
18.2663 | -106.4881
18.27 | -106.4992
18.2736 | -106.5099
18.2773 | -106.5209
18.281 | -106.532
18.2847 | -106.5406
18.2883 | -106.5489
18.292 | -106.5561
18.2957 | -106.5609
18.2994 | -106.5667
18.303 | -106.5736
18.3067 | -106.5782
18.3104 | -106.581
18.3141 | -106.5838
18.3177 | -106.5881
18.3214 | -106.5931
18.3251 | -106.598
18.3287 | -106.5979
18.3324 | -106.5929
18.3361 | -106.5914
18.3398 | -106.5916
18.3434 | -106.5897
18.3471 | -106.5852
18.3508 | -106.581
18.3545 | -106.579
18.3581 | -106.5766
18.3618 | -106.5749
18.3655 | -106.5719
18.3691 | -106.5683
18.3728 | -106.5625
18.3765 | -106.5576
18.3802 | -106.5525
18.3838 | -106.547
18.3875 | -106.5391
18.3912 | -106.5313
18.3949 | -106.5256
18.3985 | -106.5218
18.4022 | -106.517
18.4059 | -106.5124
18.4095 | -106.5096
18.4132 | -106.5041
18.4169 | -106.4982
18.4206 | -106.4948
18.4242 | -106.4933
18.4279 | -106.4936
18.4316 | -106.4939
18.4353 | -106.492
18.4389 | -106.4914
18.4426 | -106.4889
18.4463 | -106.4869
18.45 | -106.4866
18.4536 | -106.4869
18.4573 | -106.4892
18.461 | -106.4911
18.4646 | -106.4938
18.4683 | -106.4977
18.472 | -106.5022
18.4757 | -106.5087
18.4793 | -106.5166
18.483 | -106.5239
18.4867 | -106.5316
18.4904 | -106.5404
18.494 | -106.5502
18.4977 | -106.5604
18.5014 | -106.5729
18.5051 | -106.587
18.5087 | -106.6023
18.5124 | -106.6176
18.5161 | -106.6324
18.5197 | -106.6452
18.5234 | -106.6585
18.5271 | -106.6728
18.5308 | -106.6858
18.5344 | -106.6991
18.5381 | -106.7127
18.5418 | -106.7266
18.5455 | -106.7363
18.5491 | -106.7495
18.5528 | -106.7636
18.5565 | -106.7759
18.5601 | -106.7846
18.5638 | -106.7907
18.5675 | -106.7974
18.5712 | -106.8047
18.5748 | -106.8125
18.5785 | -106.8195
18.5822 | -106.825
18.5859 | -106.8284
18.5895 | -106.831
18.5932 | -106.8349
18.5969 | -106.8387
18.6006 | -106.8383
18.6042 | -106.8349
18.6079 | -106.8343
18.6116 | -106.8345
18.6152 | -106.8316
18.6189 | -106.8273
18.6226 | -106.8225
18.6263 | -106.8164
18.6299 | -106.8114
18.6336 | -106.8075
18.6373 | -106.805
18.641 | -106.8027
18.6446 | -106.8016
18.6483 | -106.7979
18.652 | -106.7945
18.6556 | -106.7906
18.6593 | -106.7862
18.663 | -106.7792
18.6667 | -106.7751
18.6703 | -106.7721
18.674 | -106.7678
18.6777 | -106.7675
18.6814 | -106.7657
18.685 | -106.7618
18.6887 | -106.7587
18.6924 | -106.7549
18.6961 | -106.7523
18.6997 | -106.7527
18.7034 | -106.7521
18.7071 | -106.752
18.7107 | -106.7511
18.7144 | -106.7507
18.7181 | -106.7509
18.7218 | -106.7516
18.7254 | -106.7522
18.7291 | -106.7516
18.7328 | -106.7512
18.7365 | -106.753
18.7401 | -106.7538
18.7438 | -106.7547
18.7475 | -106.756
18.7511 | -106.7566
18.7548 | -106.7583
18.7585 | -106.7613
18.7622 | -106.7654
18.7658 | -106.7686
18.7695 | -106.7744
18.7732 | -106.7792
18.7769 | -106.7837
18.7805 | -106.7891
18.7842 | -106.7962
18.7879 | -106.8021
18.7916 | -106.8099
18.7952 | -106.8178
18.7989 | -106.8242
18.8026 | -106.8314
18.8062 | -106.8403
18.8099 | -106.8494
18.8136 | -106.8573
18.8173 | -106.8632
18.8209 | -106.8691
18.8246 | -106.8761
18.8283 | -106.8841
18.832 | -106.8941
18.8356 | -106.8988
18.8393 | -106.9031
18.843 | -106.9075
18.8466 | -106.9099
18.8503 | -106.9104
18.854 | -106.9103
18.8577 | -106.9115
18.8613 | -106.9104
18.865 | -106.9111
18.8687 | -106.9111
18.8724 | -106.9092
18.876 | -106.9058
18.8797 | -106.9019
18.8834 | -106.8968
18.8871 | -106.8896
18.8907 | -106.8856
18.8944 | -106.8803
18.8981 | -106.8727
18.9017 | -106.8652
18.9054 | -106.858
18.9091 | -106.8493
18.9128 | -106.8384
18.9164 | -106.8299
18.9201 | -106.8228
18.9238 | -106.8153
18.9275 | -106.8074
18.9311 | -106.8004
18.9348 | -106.7918
18.9385 | -106.7839
18.9421 | -106.7774
18.9458 | -106.7717
18.9495 | -106.7677
18.9532 | -106.765
18.9568 | -106.7598
18.9605 | -106.7566
18.9642 | -106.7551
18.9679 | -106.7518
18.9715 | -106.7489
18.9752 | -106.7481
18.9789 | -106.7481
18.9826 | -106.7466
18.9862 | -106.7456
18.9899 | -106.7456
18.9936 | -106.7479
18.9972 | -106.7507
19.0009 | -106.7528
19.0046 | -106.7551
19.0083 | -106.7592
19.0119 | -106.7633
19.0156 | -106.7689
19.0193 | -106.7765
19.023 | -106.7838
19.0266 | -106.788
19.0303 | -106.7937
19.034 | -106.8011
19.0376 | -106.8082
19.0413 | -106.8166
19.045 | -106.8264
19.0487 | -106.8363
19.0523 | -106.8467
19.056 | -106.8572
19.0597 | -106.866
19.0634 | -106.876
19.067 | -106.886
19.0707 | -106.8938
19.0744 | -106.9043
19.0781 | -106.911
19.0817 | -106.9197
19.0854 | -106.9295
19.0891 | -106.9368
19.0927 | -106.9434
19.0964 | -106.9492
19.1001 | -106.9577
19.1038 | -106.9636
19.1074 | -106.9687
19.1111 | -106.9727
19.1148 | -106.9771
19.1185 | -106.9809
19.1221 | -106.984
19.1258 | -106.9851
19.1295 | -106.983
19.1332 | -106.9815
19.1368 | -106.9753
19.1405 | -106.9701
19.1442 | -106.9651
19.1478 | -106.9572
19.1515 | -106.9469
19.1552 | -106.9386
19.1589 | -106.9294
19.1625 | -106.9156
19.1662 | -106.9022
19.1699 | -106.8893
19.1736 | -106.8748
19.1772 | -106.8588
19.1809 | -106.8427
19.1846 | -106.8276
19.1882 | -106.8096
19.1919 | -106.7891
19.1956 | -106.7692
19.1993 | -106.7503
19.2029 | -106.7326
19.2066 | -106.7133
19.2103 | -106.6939
19.214 | -106.6759
19.2176 | -106.6571
19.2213 | -106.6378
19.225 | -106.6195
19.2286 | -106.605
19.2323 | -106.5912
19.236 | -106.5772
19.2397 | -106.5643
19.2433 | -106.5516
19.247 | -106.5399
19.2507 | -106.5303
19.2544 | -106.5215
19.258 | -106.5146
19.2617 | -106.5087
19.2654 | -106.5018
19.2691 | -106.4962
19.2727 | -106.4921
19.2764 | -106.4895
19.2801 | -106.4877
19.2837 | -106.4858
19.2874 | -106.4833
19.2911 | -106.4808
19.2948 | -106.4791
19.2984 | -106.4796
19.3021 | -106.4809
19.3058 | -106.4804
19.3095 | -106.4785
19.3131 | -106.4772
19.3168 | -106.4743
19.3205 | -106.4722
19.3242 | -106.4682
19.3278 | -106.4641
19.3315 | -106.4599
19.3352 | -106.4531
19.3388 | -106.4487
19.3425 | -106.4441
19.3462 | -106.4368
19.3499 | -106.4276
19.3535 | -106.4179
19.3572 | -106.4078
19.3609 | -106.3993
19.3646 | -106.3919
19.3682 | -106.3805
19.3719 | -106.3672
19.3756 | -106.3507
19.3792 | -106.3358
19.3829 | -106.3229
19.3866 | -106.3099
19.3903 | -106.2971
19.3939 | -106.2843
19.3976 | -106.271
19.4013 | -106.2556
19.405 | -106.2414
19.4086 | -106.2302
19.4123 | -106.2197
19.416 | -106.2079
19.4197 | -106.1938
19.4233 | -106.1822
19.427 | -106.1731
19.4307 | -106.1635
19.4343 | -106.1541
19.438 | -106.1453
19.4417 | -106.1373
19.4454 | -106.131
19.449 | -106.1267
19.4527 | -106.1222
19.4564 | -106.1199
19.4601 | -106.1202
19.4637 | -106.1212
19.4674 | -106.1229
19.4711 | -106.125
19.4747 | -106.1275
19.4784 | -106.1308
19.4821 | -106.1362
19.4858 | -106.1434
19.4894 | -106.1487
19.4931 | -106.1528
19.4968 | -106.1596
19.5005 | -106.1692
19.5041 | -106.1773
19.5078 | -106.1851
19.5115 | -106.1923
19.5152 | -106.1991
19.5188 | -106.2076
19.5225 | -106.2181
19.5262 | -106.2274
19.5298 | -106.2357
19.5335 | -106.2413
19.5372 | -106.2469
19.5409 | -106.2548
19.5445 | -106.2626
19.5482 | -106.2697
19.5519 | -106.2759
19.5556 | -106.2814
19.5592 | -106.2876
19.5629 | -106.293
19.5666 | -106.2976
19.5702 | -106.3033
19.5739 | -106.309
19.5776 | -106.3109
19.5813 | -106.3133
19.5849 | -106.3169
19.5886 | -106.3204
19.5923 | -106.3242
19.596 | -106.3264
19.5996 | -106.329
19.6033 | -106.3277
19.607 | -106.3303
19.6107 | -106.3308
19.6143 | -106.3318
19.618 | -106.3308
19.6217 | -106.3308
19.6253 | -106.33
19.629 | -106.3271
19.6327 | -106.3264
19.6364 | -106.3237
19.64 | -106.3203
19.6437 | -106.3162
19.6474 | -106.3142
19.6511 | -106.3132
19.6547 | -106.3115
19.6584 | -106.3097
19.6621 | -106.3073
19.6657 | -106.3066
19.6694 | -106.3038
19.6731 | -106.3009
19.6768 | -106.3003
19.6804 | -106.3036
19.6841 | -106.3047
19.6878 | -106.3037
19.6915 | -106.3051
19.6951 | -106.3075
19.6988 | -106.3123
19.7025 | -106.3173
19.7062 | -106.323
19.7098 | -106.3288
19.7135 | -106.3362
19.7172 | -106.3437
19.7208 | -106.3512
19.7245 | -106.3598
19.7282 | -106.3685
19.7319 | -106.379
19.7355 | -106.3894
19.7392 | -106.3999
19.7429 | -106.4111
19.7466 | -106.4203
19.7502 | -106.4306
19.7539 | -106.4421
19.7576 | -106.4546
19.7612 | -106.4655
19.7649 | -106.4768
19.7686 | -106.4907
19.7723 | -106.5048
19.7759 | -106.5161
19.7796 | -106.527
19.7833 | -106.5382
19.787 | -106.55
19.7906 | -106.5621
19.7943 | -106.5728
19.798 | -106.5831
19.8017 | -106.5898
19.8053 | -106.596
19.809 | -106.601
19.8127 | -106.6088
19.8163 | -106.6152
19.82 | -106.6188
19.8237 | -106.6219
19.8274 | -106.6228
19.831 | -106.6214
19.8347 | -106.6175
19.8384 | -106.6127
19.8421 | -106.6078
19.8457 | -106.6027
19.8494 | -106.595
19.8531 | -106.5856
19.8567 | -106.5789
19.8604 | -106.57
19.8641 | -106.5599
19.8678 | -106.5492
19.8714 | -106.5381
19.8751 | -106.5266
19.8788 | -106.513
19.8825 | -106.4989
19.8861 | -106.4834
19.8898 | -106.4686
19.8935 | -106.4555
19.8972 | -106.4408
19.9008 | -106.427
19.9045 | -106.4134
19.9082 | -106.4007
19.9118 | -106.3908
19.9155 | -106.3782
19.9192 | -106.3664
19.9229 | -106.3545
19.9265 | -106.3446
19.9302 | -106.3342
19.9339 | -106.3241
19.9376 | -106.3149
19.9412 | -106.3052
19.9449 | -106.2956
19.9486 | -106.2878
19.9522 | -106.2824
19.9559 | -106.278
19.9596 | -106.2735
19.9633 | -106.2666
19.9669 | -106.2618
19.9706 | -106.2569
19.9743 | -106.2524
19.978 | -106.2499
19.9816 | -106.2494
19.9853 | -106.2477
19.989 | -106.2444
19.9927 | -106.2417
19.9963 | -106.2382
20 | -106.2367
20.0037 | -106.2357
20.0073 | -106.2327
20.011 | -106.2312
20.0147 | -106.2311
20.0184 | -106.2316
20.022 | -106.2308
20.0257 | -106.2281
20.0294 | -106.2275
20.0331 | -106.2275
20.0367 | -106.2237
20.0404 | -106.2189
20.0441 | -106.2162
20.0478 | -106.2122
20.0514 | -106.2055
20.0551 | -106.1998
20.0588 | -106.1936
20.0624 | -106.1865
20.0661 | -106.18
20.0698 | -106.1711
20.0735 | -106.1625
20.0771 | -106.1553
20.0808 | -106.1485
20.0845 | -106.1413
20.0882 | -106.1345
20.0918 | -106.1281
20.0955 | -106.1194
20.0992 | -106.1109
20.1028 | -106.1031
20.1065 | -106.0963
20.1102 | -106.091
20.1139 | -106.0865
20.1175 | -106.0829
20.1212 | -106.0801
20.1249 | -106.0785
20.1286 | -106.077
20.1322 | -106.077
20.1359 | -106.0839
20.1396 | -106.0852
20.1433 | -106.0857
20.1469 | -106.0891
20.1506 | -106.0914
20.1543 | -106.0948
20.1579 | -106.0991
20.1616 | -106.1032
20.1653 | -106.1103
20.169 | -106.1191
20.1726 | -106.1277
20.1763 | -106.1353
20.18 | -106.1443
20.1837 | -106.154
20.1873 | -106.1627
20.191 | -106.1721
20.1947 | -106.1822
20.1983 | -106.1894
20.202 | -106.1968
20.2057 | -106.204
20.2094 | -106.209
20.213 | -106.2125
20.2167 | -106.2163
20.2204 | -106.219
20.2241 | -106.2214
20.2277 | -106.2232
20.2314 | -106.2269
20.2351 | -106.2265
20.2388 | -106.2236
20.2424 | -106.2204
20.2461 | -106.2169
20.2498 | -106.2153
20.2534 | -106.2137
20.2571 | -106.2101
20.2608 | -106.2041
20.2645 | -106.1976
20.2681 | -106.1903
20.2718 | -106.1829
20.2755 | -106.1736
20.2792 | -106.1635
20.2828 | -106.154
20.2865 | -106.1433
20.2902 | -106.1318
20.2938 | -106.1226
20.2975 | -106.112
20.3012 | -106.0991
20.3049 | -106.0855
20.3085 | -106.0737
20.3122 | -106.0643
20.3159 | -106.0534
20.3196 | -106.0411
20.3232 | -106.0324
20.3269 | -106.0227
20.3306 | -106.012
20.3343 | -106.002
20.3379 | -105.9943
20.3416 | -105.9875
20.3453 | -105.9806
20.3489 | -105.9716
20.3526 | -105.9618
20.3563 | -105.9547
20.36 | -105.9488
20.3636 | -105.9427
20.3673 | -105.9365
20.371 | -105.9283
20.3747 | -105.9199
20.3783 | -105.9143
20.382 | -105.9093
20.3857 | -105.9039
20.3893 | -105.9
20.393 | -105.894
20.3967 | -105.8885
20.4004 | -105.8849
20.404 | -105.8823
20.4077 | -105.8776
20.4114 | -105.8708
20.4151 | -105.8635
20.4187 | -105.8581
20.4224 | -105.854
20.4261 | -105.8467
20.4298 | -105.8394
20.4334 | -105.8337
20.4371 | -105.8258
20.4408 | -105.8209
20.4444 | -105.8146
20.4481 | -105.8073
20.4518 | -105.8009
20.4555 | -105.7929
20.4591 | -105.7835
20.4628 | -105.7736
20.4665 | -105.7649
20.4702 | -105.7531
20.4738 | -105.7428
20.4775 | -105.7326
20.4812 | -105.7203
20.4848 | -105.7105
20.4885 | -105.7014
20.4922 | -105.6925
20.4959 | -105.682
20.4995 | -105.6723
20.5032 | -105.663
20.5069 | -105.6528
20.5106 | -105.6433
20.5142 | -105.6343
20.5179 | -105.6247
20.5216 | -105.6154
20.5253 | -105.608
20.5289 | -105.6009
20.5326 | -105.5962
20.5363 | -105.5924
20.5399 | -105.5876
20.5436 | -105.5818
20.5473 | -105.58
20.551 | -105.578
20.5546 | -105.5756
20.5583 | -105.574
20.562 | -105.5739
20.5657 | -105.5775
20.5693 | -105.5817
20.573 | -105.5858
20.5767 | -105.5908
20.5803 | -105.595
20.584 | -105.598
20.5877 | -105.6022
20.5914 | -105.6088
20.595 | -105.6155
20.5987 | -105.6217
20.6024 | -105.6271
20.6061 | -105.6334
20.6097 | -105.6365
20.6134 | -105.6388
20.6171 | -105.6439
20.6208 | -105.6491
20.6244 | -105.6535
20.6281 | -105.6562
20.6318 | -105.6563
20.6354 | -105.6566
20.6391 | -105.6596
20.6428 | -105.658
20.6465 | -105.6564
20.6501 | -105.6553
20.6538 | -105.6536
20.6575 | -105.6486
20.6612 | -105.6451
20.6648 | -105.6353
20.6685 | -105.6313
20.6722 | -105.6235
20.6758 | -105.6155
20.6795 | -105.6071
20.6832 | -105.5963
20.6869 | -105.5863
20.6905 | -105.5759
20.6942 | -105.5609
20.6979 | -105.5472
20.7016 | -105.5357
20.7052 | -105.524
20.7089 | -105.5127
20.7126 | -105.4988
20.7163 | -105.4831
20.7199 | -105.4703
20.7236 | -105.4562
20.7273 | -105.4399
20.7309 | -105.4257
20.7346 | -105.4095
20.7383 | -105.394
20.742 | -105.3802
20.7456 | -105.3671
20.7493 | -105.354
20.753 | -105.3422
20.7567 | -105.3307
20.7603 | -105.3202
20.764 | -105.3131
20.7677 | -105.3057
20.7714 | -105.2974
20.775 | -105.2922
20.7787 | -105.287
20.7824 | -105.2831
20.786 | -105.2801
20.7897 | -105.2796
20.7934 | -105.2801
20.7971 | -105.2812
20.8007 | -105.2849
20.8044 | -105.289
20.8081 | -105.2947
20.8118 | -105.3037
20.8154 | -105.3129
20.8191 | -105.3209
20.8228 | -105.3314
20.8264 | -105.3414
20.8301 | -105.3516
20.8338 | -105.3637
20.8375 | -105.3789
20.8411 | -105.3921
20.8448 | -105.4074
20.8485 | -105.4238
20.8522 | -105.4407
20.8558 | -105.4539
20.8595 | -105.4654
20.8632 | -105.4798
20.8668 | -105.4946
20.8705 | -105.5077
20.8742 | -105.52
20.8779 | -105.5333
20.8815 | -105.5463
20.8852 | -105.5586
20.8889 | -105.5708
20.8926 | -105.5827
20.8962 | -105.5937
20.8999 | -105.6035
20.9036 | -105.6144
20.9073 | -105.6226
20.9109 | -105.6268
20.9146 | -105.6311
20.9183 | -105.6343
20.9219 | -105.6359
20.9256 | -105.6383
20.9293 | -105.6403
20.933 | -105.6382
20.9366 | -105.6387
20.9403 | -105.635
20.944 | -105.6291
20.9477 | -105.6221
20.9513 | -105.6141
20.955 | -105.6057
20.9587 | -105.5966
20.9624 | -105.5854
20.966 | -105.5747
20.9697 | -105.5646
20.9734 | -105.554
20.977 | -105.5424
20.9807 | -105.5331
20.9844 | -105.5226
20.9881 | -105.5103
20.9917 | -105.4998
20.9954 | -105.4873
20.9991 | -105.4743
21.0028 | -105.4624
21.0064 | -105.4495
21.0101 | -105.4364
21.0138 | -105.4251
21.0174 | -105.4143
21.0211 | -105.4063
21.0248 | -105.3986
21.0285 | -105.3885
21.0321 | -105.3792
21.0358 | -105.3708
21.0395 | -105.3634
21.0432 | -105.3564
21.0468 | -105.3493
21.0505 | -105.3413
21.0542 | -105.3368
21.0579 | -105.332
21.0615 | -105.326
21.0652 | -105.3208
21.0689 | -105.3199
21.0725 | -105.3184
21.0762 | -105.3158
21.0799 | -105.3134
21.0836 | -105.312
21.0872 | -105.3124
21.0909 | -105.3141
21.0946 | -105.3158
21.0983 | -105.3171
21.1019 | -105.3172
21.1056 | -105.3169
21.1093 | -105.3168
21.1129 | -105.3166
21.1166 | -105.3188
21.1203 | -105.3222
21.124 | -105.3263
21.1276 | -105.329
21.1313 | -105.3293
21.135 | -105.3296
21.1387 | -105.3311
21.1423 | -105.3329
21.146 | -105.3337
21.1497 | -105.3341
21.1534 | -105.3389
21.157 | -105.3432
21.1607 | -105.3432
21.1644 | -105.3421
21.168 | -105.3414
21.1717 | -105.3418
21.1754 | -105.3386
21.1791 | -105.3365
21.1827 | -105.3348
21.1864 | -105.3319
21.1901 | -105.3294
21.1938 | -105.3319
21.1974 | -105.3321
21.2011 | -105.3325
21.2048 | -105.3318
21.2084 | -105.3312
21.2121 | -105.3325
21.2158 | -105.3329
21.2195 | -105.3332
21.2231 | -105.3333
21.2268 | -105.3338
21.2305 | -105.3347
21.2342 | -105.3365
21.2378 | -105.3401
21.2415 | -105.3446
21.2452 | -105.3491
21.2489 | -105.3549
21.2525 | -105.3613
21.2562 | -105.3659
21.2599 | -105.37
21.2635 | -105.376
21.2672 | -105.3821
21.2709 | -105.3883
21.2746 | -105.393
21.2782 | -105.3995
21.2819 | -105.4055
21.2856 | -105.4128
21.2893 | -105.4201
21.2929 | -105.4275
21.2966 | -105.4364
21.3003 | -105.4447
21.3039 | -105.4528
21.3076 | -105.4619
21.3113 | -105.4715
21.315 | -105.4809
21.3186 | -105.4896
21.3223 | -105.4955
21.326 | -105.4988
21.3297 | -105.5053
21.3333 | -105.5133
21.337 | -105.5187
21.3407 | -105.5244
21.3444 | -105.5303
21.348 | -105.5358
21.3517 | -105.5425
21.3554 | -105.5459
21.359 | -105.5479
21.3627 | -105.5496
21.3664 | -105.552
21.3701 | -105.5526
21.3737 | -105.5524
21.3774 | -105.5525
21.3811 | -105.5519
21.3848 | -105.5508
21.3884 | -105.5473
21.3921 | -105.5411
21.3958 | -105.5338
21.3994 | -105.5277
21.4031 | -105.5248
21.4068 | -105.5215
21.4105 | -105.5159
21.4141 | -105.5106
21.4178 | -105.5058
21.4215 | -105.5005
21.4252 | -105.4944
21.4288 | -105.4891
21.4325 | -105.4844
21.4362 | -105.4785
21.4399 | -105.4716
21.4435 | -105.4641
21.4472 | -105.4565
21.4509 | -105.4493
21.4545 | -105.4406
21.4582 | -105.4317
21.4619 | -105.4252
21.4656 | -105.419
21.4692 | -105.4133
21.4729 | -105.4084
21.4766 | -105.4038
21.4803 | -105.398
21.4839 | -105.3921
21.4876 | -105.3883
21.4913 | -105.3836
21.4949 | -105.3768
21.4986 | -105.3695
21.5023 | -105.3638
21.506 | -105.3597
21.5096 | -105.3591
21.5133 | -105.3586
21.517 | -105.3563
21.5207 | -105.353
21.5243 | -105.3503
21.528 | -105.3493
21.5317 | -105.3475
21.5354 | -105.3439
21.539 | -105.3423
21.5427 | -105.3406
21.5464 | -105.3388
21.55 | -105.3379
21.5537 | -105.3376
21.5574 | -105.3408
21.5611 | -105.3438
21.5647 | -105.3427
21.5684 | -105.3432
21.5721 | -105.3444
21.5758 | -105.3441
21.5794 | -105.3454
21.5831 | -105.3475
21.5868 | -105.3502
21.5905 | -105.3517
21.5941 | -105.3536
21.5978 | -105.3566
21.6015 | -105.3589
21.6051 | -105.3627
21.6088 | -105.3662
21.6125 | -105.3702
21.6162 | -105.3759
21.6198 | -105.3828
21.6235 | -105.3892
21.6272 | -105.3949
21.6309 | -105.4038
21.6345 | -105.4105
21.6382 | -105.4156
21.6419 | -105.4229
21.6455 | -105.4289
21.6492 | -105.4342
21.6529 | -105.4408
21.6566 | -105.4467
21.6602 | -105.4522
21.6639 | -105.4584
21.6676 | -105.4653
21.6713 | -105.4704
21.6749 | -105.4739
21.6786 | -105.4779
21.6823 | -105.4823
21.6859 | -105.4873
21.6896 | -105.491
21.6933 | -105.4912
21.697 | -105.4942
21.7006 | -105.4981
21.7043 | -105.4984
21.708 | -105.4978
21.7117 | -105.4975
21.7153 | -105.4953
21.719 | -105.4883
21.7227 | -105.4811
21.7264 | -105.4722
21.73 | -105.466
21.7337 | -105.4594
21.7374 | -105.4532
21.741 | -105.4457
21.7447 | -105.4363
21.7484 | -105.4272
21.7521 | -105.4183
21.7557 | -105.4099
21.7594 | -105.401
21.7631 | -105.3906
21.7668 | -105.3803
21.7704 | -105.3686
21.7741 | -105.3564
21.7778 | -105.3461
21.7815 | -105.3354
21.7851 | -105.3256
21.7888 | -105.3139
21.7925 | -105.3008
21.7961 | -105.2895
21.7998 | -105.2774
21.8035 | -105.2662
21.8072 | -105.2575
21.8108 | -105.2492
21.8145 | -105.2399
21.8182 | -105.2284
21.8219 | -105.2196
21.8255 | -105.2101
21.8292 | -105.2025
21.8329 | -105.1983
21.8365 | -105.1935
21.8402 | -105.188
21.8439 | -105.1837
21.8476 | -105.1778
21.8512 | -105.1713
21.8549 | -105.1668
21.8586 | -105.1632
21.8623 | -105.1589
21.8659 | -105.1539
21.8696 | -105.15
21.8733 | -105.1489
21.877 | -105.1478
21.8806 | -105.147
21.8843 | -105.1478
21.888 | -105.1483
21.8916 | -105.1465
21.8953 | -105.1441
21.899 | -105.1455
21.9027 | -105.148
21.9063 | -105.15
21.91 | -105.15
21.9137 | -105.1501
21.9174 | -105.1516
21.921 | -105.1536
21.9247 | -105.1532
21.9284 | -105.1542
21.932 | -105.1568
21.9357 | -105.1589
21.9394 | -105.1589
21.9431 | -105.1603
21.9467 | -105.1629
21.9504 | -105.1638
21.9541 | -105.1635
21.9578 | -105.1642
21.9614 | -105.166
21.9651 | -105.1674
21.9688 | -105.1671
21.9725 | -105.167
21.9761 | -105.1677
21.9798 | -105.1689
21.9835 | -105.1697
21.9871 | -105.169
21.9908 | -105.1683
21.9945 | -105.1678
21.9982 | -105.1685
22.0018 | -105.1709
22.0055 | -105.1718
22.0092 | -105.1692
22.0129 | -105.1679
22.0165 | -105.1652
22.0202 | -105.163
22.0239 | -105.1646
22.0275 | -105.1691
22.0312 | -105.1723
22.0349 | -105.1735
22.0386 | -105.1737
22.0422 | -105.1748
22.0459 | -105.1749
22.0496 | -105.174
22.0533 | -105.1769
22.0569 | -105.1809
22.0606 | -105.1848
22.0643 | -105.1885
22.068 | -105.1911
22.0716 | -105.1944
22.0753 | -105.1993
22.079 | -105.2045
22.0826 | -105.2097
22.0863 | -105.2145
22.09 | -105.2171
22.0937 | -105.2194
22.0973 | -105.2233
22.101 | -105.2276
22.1047 | -105.232
22.1084 | -105.2353
22.112 | -105.24
22.1157 | -105.2447
22.1194 | -105.2496
22.123 | -105.2537
22.1267 | -105.2575
22.1304 | -105.2625
22.1341 | -105.2658
22.1377 | -105.2713
22.1414 | -105.2764
22.1451 | -105.28
22.1488 | -105.285
22.1524 | -105.2891
22.1561 | -105.2932
22.1598 | -105.2998
22.1635 | -105.3068
22.1671 | -105.3143
22.1708 | -105.3212
22.1745 | -105.3276
22.1781 | -105.3344
22.1818 | -105.342
22.1855 | -105.3506
22.1892 | -105.3585
22.1928 | -105.3655
22.1965 | -105.3728
22.2002 | -105.3807
22.2039 | -105.3894
22.2075 | -105.3985
22.2112 | -105.4063
22.2149 | -105.4144
22.2185 | -105.4229
22.2222 | -105.4335
22.2259 | -105.4422
22.2296 | -105.4494
22.2332 | -105.4565
22.2369 | -105.4661
22.2406 | -105.4746
22.2443 | -105.4822
22.2479 | -105.4917
22.2516 | -105.5009
22.2553 | -105.508
22.259 | -105.5163
22.2626 | -105.524
22.2663 | -105.5313
22.27 | -105.5374
22.2736 | -105.5435
22.2773 | -105.5484
22.281 | -105.5526
22.2847 | -105.5591
22.2883 | -105.5652
22.292 | -105.5698
22.2957 | -105.5732
22.2994 | -105.5782
22.303 | -105.5835
22.3067 | -105.5877
22.3104 | -105.5911
22.3141 | -105.594
22.3177 | -105.5967
22.3214 | -105.5991
22.3251 | -105.6035
22.3287 | -105.6073
22.3324 | -105.6109
22.3361 | -105.6137
22.3398 | -105.6133
22.3434 | -105.6138
22.3471 | -105.6142
22.3508 | -105.6152
22.3545 | -105.6159
22.3581 | -105.6173
22.3618 | -105.6197
22.3655 | -105.6229
22.3691 | -105.6222
22.3728 | -105.6222
22.3765 | -105.6232
22.3802 | -105.6222
22.3838 | -105.6213
22.3875 | -105.6151
22.3912 | -105.6072
22.3949 | -105.5971
22.3985 | -105.588
22.4022 | -105.5796
22.4059 | -105.5696
22.4095 | -105.5577
22.4132 | -105.5442
22.4169 | -105.5309
22.4206 | -105.5178
22.4242 | -105.504
22.4279 | -105.4891
22.4316 | -105.4714
22.4353 | -105.4514
22.4389 | -105.4314
22.4426 | -105.4109
22.4463 | -105.3895
22.45 | -105.3676
22.4536 | -105.3432
22.4573 | -105.317
22.461 | -105.2911
22.4646 | -105.2652
22.4683 | -105.2402
22.472 | -105.2134
22.4757 | -105.1848
22.4793 | -105.1563
22.483 | -105.1275
22.4867 | -105.0958
22.4904 | -105.0661
22.494 | -105.0382
22.4977 | -105.0099
22.5014 | -104.9811
22.5051 | -104.9536
22.5087 | -104.9276
22.5124 | -104.9018
22.5161 | -104.8775
22.5197 | -104.8539
22.5234 | -104.8307
22.5271 | -104.8071
22.5308 | -104.7865
22.5344 | -104.7685
22.5381 | -104.7531
22.5418 | -104.7399
22.5455 | -104.7299
22.5491 | -104.7213
22.5528 | -104.7129
22.5565 | -104.706
22.5601 | -104.7043
22.5638 | -104.7035
22.5675 | -104.7059
22.5712 | -104.71
22.5748 | -104.7126
22.5785 | -104.7191
22.5822 | -104.7282
22.5859 | -104.7401
22.5895 | -104.7542
22.5932 | -104.769
22.5969 | -104.7857
22.6006 | -104.8055
22.6042 | -104.8257
22.6079 | -104.8451
22.6116 | -104.864
22.6152 | -104.8842
22.6189 | -104.9058
22.6226 | -104.929
22.6263 | -104.9503
22.6299 | -104.9694
22.6336 | -104.9892
22.6373 | -105.0103
22.641 | -105.0311
22.6446 | -105.0521
22.6483 | -105.0711
22.652 | -105.0886
22.6556 | -105.1057
22.6593 | -105.122
22.663 | -105.1358
22.6667 | -105.1489
22.6703 | -105.1612
22.674 | -105.1733
22.6777 | -105.1829
22.6814 | -105.1929
22.685 | -105.2023
22.6887 | -105.2112
22.6924 | -105.2184
22.6961 | -105.2244
22.6997 | -105.2313
22.7034 | -105.2384
22.7071 | -105.2484
22.7107 | -105.2594
22.7144 | -105.2727
22.7181 | -105.2829
22.7218 | -105.2916
22.7254 | -105.301
22.7291 | -105.3103
22.7328 | -105.3192
22.7365 | -105.3285
22.7401 | -105.3363
22.7438 | -105.3427
22.7475 | -105.3502
22.7511 | -105.3589
22.7548 | -105.3675
22.7585 | -105.3755
22.7622 | -105.3847
22.7658 | -105.3941
22.7695 | -105.4041
22.7732 | -105.4166
22.7769 | -105.4297
22.7805 | -105.4397
22.7842 | -105.4513
22.7879 | -105.4599
22.7916 | -105.4695
22.7952 | -105.4787
22.7989 | -105.4869
22.8026 | -105.4968
22.8062 | -105.5066
22.8099 | -105.5185
22.8136 | -105.5286
22.8173 | -105.536
22.8209 | -105.5438
22.8246 | -105.5531
22.8283 | -105.56
22.832 | -105.565
22.8356 | -105.571
22.8393 | -105.5746
22.843 | -105.5785
22.8466 | -105.5831
22.8503 | -105.5892
22.854 | -105.5909
22.8577 | -105.5917
22.8613 | -105.59
22.865 | -105.5884
22.8687 | -105.5867
22.8724 | -105.584
22.876 | -105.582
22.8797 | -105.5802
22.8834 | -105.5747
22.8871 | -105.5697
22.8907 | -105.566
22.8944 | -105.5593
22.8981 | -105.5548
22.9017 | -105.5495
22.9054 | -105.5421
22.9091 | -105.5355
22.9128 | -105.5303
22.9164 | -105.5229
22.9201 | -105.5135
22.9238 | -105.5031
22.9275 | -105.4929
22.9311 | -105.4837
22.9348 | -105.4756
22.9385 | -105.4689
22.9421 | -105.4626
22.9458 | -105.4555
22.9495 | -105.4489
22.9532 | -105.4428
22.9568 | -105.4354
22.9605 | -105.4272
22.9642 | -105.4178
22.9679 | -105.4107
22.9715 | -105.4041
22.9752 | -105.397
22.9789 | -105.3897
22.9826 | -105.3839
22.9862 | -105.3798
22.9899 | -105.3752
22.9936 | -105.3688
22.9972 | -105.3624
23.0009 | -105.3579
23.0046 | -105.3525
23.0083 | -105.3477
23.0119 | -105.3434
23.0156 | -105.3379
23.0193 | -105.3321
23.023 | -105.3284
23.0266 | -105.3236
23.0303 | -105.3176
23.034 | -105.3101
23.0376 | -105.3019
23.0413 | -105.2965
23.045 | -105.292
23.0487 | -105.2879
23.0523 | -105.2844
23.056 | -105.2805
23.0597 | -105.2763
23.0634 | -105.2741
23.067 | -105.2725
23.0707 | -105.2684
23.0744 | -105.2652
23.0781 | -105.2642
23.0817 | -105.2637
23.0854 | -105.262
23.0891 | -105.2596
23.0927 | -105.2587
23.0964 | -105.2575
23.1001 | -105.2563
23.1038 | -105.257
23.1074 | -105.2581
23.1111 | -105.2593
23.1148 | -105.26
23.1185 | -105.2621
23.1221 | -105.2672
23.1258 | -105.2728
23.1295 | -105.277
23.1332 | -105.2799
23.1368 | -105.2833
23.1405 | -105.2875
23.1442 | -105.2921
23.1478 | -105.2967
23.1515 | -105.2994
23.1552 | -105.3048
23.1589 | -105.3121
23.1625 | -105.3191
23.1662 | -105.3267
23.1699 | -105.3337
23.1736 | -105.3405
23.1772 | -105.348
23.1809 | -105.3565
23.1846 | -105.3628
23.1882 | -105.3692
23.1919 | -105.3758
23.1956 | -105.3807
23.1993 | -105.3855
23.2029 | -105.3912
23.2066 | -105.396
23.2103 | -105.401
23.214 | -105.4053
23.2176 | -105.4107
23.2213 | -105.4185
23.225 | -105.4241
23.2286 | -105.4285
23.2323 | -105.4346
23.236 | -105.439
23.2397 | -105.4422
23.2433 | -105.4485
23.247 | -105.4562
23.2507 | -105.4613
23.2544 | -105.4639
23.258 | -105.4676
23.2617 | -105.472
23.2654 | -105.4757
23.2691 | -105.4791
23.2727 | -105.4826
23.2764 | -105.4877
23.2801 | -105.4924
23.2837 | -105.4973
23.2874 | -105.5008
23.2911 | -105.5052
23.2948 | -105.5086
23.2984 | -105.5116
23.3021 | -105.5158
23.3058 | -105.5205
23.3095 | -105.5249
23.3131 | -105.5275
23.3168 | -105.5305
23.3205 | -105.5322
23.3242 | -105.5344
23.3278 | -105.5383
23.3315 | -105.5411
23.3352 | -105.5418
23.3388 | -105.5425
23.3425 | -105.545
23.3462 | -105.5485
23.3499 | -105.5509
23.3535 | -105.5492
23.3572 | -105.5498
23.3609 | -105.5529
23.3646 | -105.5553
23.3682 | -105.558
23.3719 | -105.558
23.3756 | -105.557
23.3792 | -105.5564
23.3829 | -105.5564
23.3866 | -105.5574
23.3903 | -105.5581
23.3939 | -105.5578
23.3976 | -105.5563
23.4013 | -105.5579
23.405 | -105.5593
23.4086 | -105.5609
23.4123 | -105.5605
23.416 | -105.5611
23.4197 | -105.5628
23.4233 | -105.5631
23.427 | -105.5622
23.4307 | -105.5623
23.4343 | -105.5632
23.438 | -105.5646
23.4417 | -105.5667
23.4454 | -105.5674
23.449 | -105.5665
23.4527 | -105.5663
23.4564 | -105.5649
23.4601 | -105.5633
23.4637 | -105.5638
23.4674 | -105.5647
23.4711 | -105.5662
23.4747 | -105.5691
23.4784 | -105.571
23.4821 | -105.5721
23.4858 | -105.5706
23.4894 | -105.571
23.4931 | -105.5725
23.4968 | -105.5739
23.5005 | -105.5759
23.5041 | -105.5762
23.5078 | -105.5764
23.5115 | -105.5784
23.5152 | -105.5815
23.5188 | -105.5858
23.5225 | -105.5901
23.5262 | -105.5956
23.5298 | -105.6015
23.5335 | -105.6046
23.5372 | -105.6088
23.5409 | -105.6129
23.5445 | -105.6186
23.5482 | -105.6216
23.5519 | -105.6248
23.5556 | -105.629
23.5592 | -105.635
23.5629 | -105.6402
23.5666 | -105.6463
23.5702 | -105.6526
23.5739 | -105.6573
23.5776 | -105.6608
23.5813 | -105.6663
23.5849 | -105.6735
23.5886 | -105.6798
23.5923 | -105.6857
23.596 | -105.6904
23.5996 | -105.6924
23.6033 | -105.6949
23.607 | -105.6983
23.6107 | -105.7033
23.6143 | -105.7057
23.618 | -105.7086
23.6217 | -105.7111
23.6253 | -105.7134
23.629 | -105.7171
23.6327 | -105.7193
23.6364 | -105.7196
23.64 | -105.7208
23.6437 | -105.7243
23.6474 | -105.7249
23.6511 | -105.7265
23.6547 | -105.7296
23.6584 | -105.732
23.6621 | -105.731
23.6657 | -105.7305
23.6694 | -105.7327
23.6731 | -105.7335
23.6768 | -105.7333
23.6804 | -105.7338
23.6841 | -105.7353
23.6878 | -105.7375
23.6915 | -105.7378
23.6951 | -105.7362
23.6988 | -105.7349
23.7025 | -105.7318
23.7062 | -105.731
23.7098 | -105.7326
23.7135 | -105.733
23.7172 | -105.7334
23.7208 | -105.7331
23.7245 | -105.7336
23.7282 | -105.7336
23.7319 | -105.7331
23.7355 | -105.7319
23.7392 | -105.731
23.7429 | -105.7303
23.7466 | -105.7304
23.7502 | -105.73
23.7539 | -105.7297
23.7576 | -105.7298
23.7612 | -105.7315
23.7649 | -105.7337
23.7686 | -105.7338
23.7723 | -105.7347
23.7759 | -105.7364
23.7796 | -105.7392
23.7833 | -105.742
23.787 | -105.7434
23.7906 | -105.7458
23.7943 | -105.7488
23.798 | -105.7523
23.8017 | -105.7548
23.8053 | -105.7568
23.809 | -105.7576
23.8127 | -105.7609
23.8163 | -105.7655
23.82 | -105.7694
23.8237 | -105.7711
23.8274 | -105.7721
23.831 | -105.7734
23.8347 | -105.7752
23.8384 | -105.7794
23.8421 | -105.7838
23.8457 | -105.7848
23.8494 | -105.7873
23.8531 | -105.7894
23.8567 | -105.7911
23.8604 | -105.7931
23.8641 | -105.7953
23.8678 | -105.7963
23.8714 | -105.7994
23.8751 | -105.802
23.8788 | -105.8042
23.8825 | -105.8067
23.8861 | -105.8088
23.8898 | -105.812
23.8935 | -105.8151
23.8972 | -105.8164
23.9008 | -105.818
23.9045 | -105.8205
23.9082 | -105.8228
23.9118 | -105.8236
23.9155 | -105.825
23.9192 | -105.8258
23.9229 | -105.8283
23.9265 | -105.8294
23.9302 | -105.8296
23.9339 | -105.829
23.9376 | -105.8303
23.9412 | -105.8324
23.9449 | -105.8336
23.9486 | -105.8338
23.9522 | -105.8342
23.9559 | -105.8333
23.9596 | -105.8324
23.9633 | -105.8305
23.9669 | -105.8288
23.9706 | -105.8273
23.9743 | -105.8266
23.978 | -105.8252
23.9816 | -105.8225
23.9853 | -105.8224
23.989 | -105.8218
23.9927 | -105.821
23.9963 | -105.8216
24 | -105.8227
24.0037 | -105.8232
24.0073 | -105.823
24.011 | -105.8234
24.0147 | -105.8235
24.0184 | -105.8237
24.022 | -105.8258
24.0257 | -105.8289
24.0294 | -105.8317
24.0331 | -105.8353
24.0367 | -105.8392
24.0404 | -105.8427
24.0441 | -105.846
24.0478 | -105.8472
24.0514 | -105.849
24.0551 | -105.8522
24.0588 | -105.856
24.0624 | -105.8592
24.0661 | -105.8636
24.0698 | -105.8697
24.0735 | -105.8756
24.0771 | -105.8809
24.0808 | -105.8848
24.0845 | -105.8907
24.0882 | -105.8965
24.0918 | -105.9025
24.0955 | -105.9101
24.0992 | -105.9169
24.1028 | -105.9217
24.1065 | -105.9269
24.1102 | -105.9317
24.1139 | -105.9359
24.1175 | -105.9422
24.1212 | -105.9476
24.1249 | -105.9515
24.1286 | -105.9557
24.1322 | -105.9585
24.1359 | -105.9622
24.1396 | -105.9676
24.1433 | -105.9733
24.1469 | -105.9783
24.1506 | -105.9821
24.1543 | -105.9859
24.1579 | -105.9908
24.1616 | -105.996
24.1653 | -105.9992
24.169 | -106.002
24.1726 | -106.0042
24.1763 | -106.0053
24.18 | -106.0062
24.1837 | -106.0074
24.1873 | -106.0083
24.191 | -106.0079
24.1947 | -106.0088
24.1983 | -106.0096
24.202 | -106.0083
24.2057 | -106.0077
24.2094 | -106.0093
24.213 | -106.0098
24.2167 | -106.0082
24.2204 | -106.0086
24.2241 | -106.0077
24.2277 | -106.0074
24.2314 | -106.0075
24.2351 | -106.0073
24.2388 | -106.0046
24.2424 | -105.9986
24.2461 | -105.9857
24.2498 | -105.9598
24.2534 | -105.9233
24.2571 | -105.8713
24.2608 | -105.7925
24.2645 | -105.6469
24.2681 | -105.4536
24.2718 | -105.2612
24.2755 | -105.1032
24.2792 | -104.9775
24.2828 | -104.8702
24.2865 | -104.7592
24.2902 | -104.6419
24.2938 | -104.5187
24.2975 | -104.3913
24.3012 | -104.2585
24.3049 | -104.1201
24.3085 | -103.9776
24.3122 | -103.8316
24.3159 | -103.682
24.3196 | -103.5292
24.3232 | -103.3745
24.3269 | -103.2164
24.3306 | -103.0567
24.3343 | -102.8973
24.3379 | -102.7369
24.3416 | -102.5776
24.3453 | -102.4199
24.3489 | -102.2645
24.3526 | -102.1115
24.3563 | -101.961
24.36 | -101.8138
24.3636 | -101.6703
24.3673 | -101.5314
24.371 | -101.3975
24.3747 | -101.2872
24.3783 | -101.1702
24.382 | -101.0588
24.3857 | -100.9537
24.3893 | -100.8551
24.393 | -100.7633
24.3967 | -100.6781
24.4004 | -100.5994
24.404 | -100.5276
24.4077 | -100.4635
24.4114 | -100.4067
24.4151 | -100.3576
24.4187 | -100.3165
24.4224 | -100.2829
24.4261 | -100.2571
24.4298 | -100.2392
24.4334 | -100.2292
24.4371 | -100.2272
24.4408 | -100.2329
24.4444 | -100.2469
24.4481 | -100.269
24.4518 | -100.2988
24.4555 | -100.3363
24.4591 | -100.3818
24.4628 | -100.4345
24.4665 | -100.4951
24.4702 | -100.5632
24.4738 | -100.6384
24.4775 | -100.7208
24.4812 | -100.8094
24.4848 | -100.9044
24.4885 | -101.006
24.4922 | -101.1137
24.4959 | -101.2273
24.4995 | -101.3469
24.5032 | -101.4721
24.5069 | -101.602
24.5106 | -101.737
24.5142 | -101.8759
24.5179 | -102.0183
24.5216 | -102.1637
24.5253 | -102.312
24.5289 | -102.4636
24.5326 | -102.6171
24.5363 | -102.772
24.5399 | -102.9268
24.5436 | -103.0816
24.5473 | -103.2371
24.551 | -103.3904
24.5546 | -103.5398
24.5583 | -103.685
24.562 | -103.8244
24.5657 | -103.9557
24.5693 | -104.074
24.573 | -104.1705
24.5767 | -104.2523
24.5803 | -104.3286
24.584 | -104.4077
24.5877 | -104.4892
24.5914 | -104.5669
24.595 | -104.644
24.5987 | -104.7153
24.6024 | -104.7803
24.6061 | -104.8473
24.6097 | -104.9421
24.6134 | -105.0458
24.6171 | -105.1151
24.6208 | -105.1281
24.6244 | -105.0912
24.6281 | -105.0233
24.6318 | -104.9517
24.6354 | -104.8777
24.6391 | -104.8017
24.6428 | -104.722
24.6465 | -104.6394
24.6501 | -104.5559
24.6538 | -104.4696
24.6575 | -104.3834
24.6612 | -104.2979
24.6648 | -104.2113
24.6685 | -104.124
24.6722 | -104.0382
24.6758 | -103.9542
24.6795 | -103.8715
24.6832 | -103.7918
24.6869 | -103.714
24.6905 | -103.6377
24.6942 | -103.5638
24.6979 | -103.4925
24.7016 | -103.425
24.7052 | -103.3606
24.7089 | -103.2997
24.7126 | -103.2422
24.7163 | -103.189
24.7199 | -103.1397
24.7236 | -103.0948
24.7273 | -103.055
24.7309 | -103.0206
24.7346 | -102.9899
24.7383 | -102.9639
24.742 | -102.9438
24.7456 | -102.9289
24.7493 | -102.9194
24.753 | -102.9152
24.7567 | -102.9156
24.7603 | -102.9214
24.764 | -102.9324
24.7677 | -102.9483
24.7714 | -102.9686
24.775 | -102.9936
24.7787 | -103.0239
24.7824 | -103.0595
24.786 | -103.0995
24.7897 | -103.1433
24.7934 | -103.1914
24.7971 | -103.2441
24.8007 | -103.3005
24.8044 | -103.3594
24.8081 | -103.4213
24.8118 | -103.4872
24.8154 | -103.5566
24.8191 | -103.6287
24.8228 | -103.7032
24.8264 | -103.779
24.8301 | -103.8572
24.8338 | -103.9369
24.8375 | -104.0166
24.8411 | -104.0971
24.8448 | -104.1781
24.8485 | -104.2589
24.8522 | -104.3383
24.8558 | -104.4184
24.8595 | -104.497
24.8632 | -104.5741
24.8668 | -104.6493
24.8705 | -104.7237
24.8742 | -104.7969
24.8779 | -104.868
24.8815 | -104.9346
24.8852 | -104.9987
24.8889 | -105.0616
24.8926 | -105.1204
24.8962 | -105.1755
24.8999 | -105.2282
24.9036 | -105.2486
24.9073 | -105.2994
24.9109 | -105.3507
24.9146 | -105.4028
24.9183 | -105.4532
24.9219 | -105.5084
24.9256 | -105.5767
24.9293 | -105.642
24.933 | -105.6939
24.9366 | -105.7255
24.9403 | -105.7388
24.944 | -105.7475
24.9477 | -105.7529
24.9513 | -105.7592
24.955 | -105.7661
24.9587 | -105.7724
24.9624 | -105.7783
24.966 | -105.7861
24.9697 | -105.7954
24.9734 | -105.799
24.977 | -105.8004
24.9807 | -105.7994
24.9844 | -105.7966
24.9881 | -105.7954
24.9917 | -105.7941
24.9954 | -105.7923
24.9991 | -105.7912
25.0028 | -105.7893
25.0064 | -105.7886
25.0101 | -105.7876
25.0138 | -105.7846
25.0174 | -105.7839
25.0211 | -105.7854
25.0248 | -105.7874
25.0285 | -105.7889
25.0321 | -105.7891
25.0358 | -105.7901
25.0395 | -105.7912
25.0432 | -105.7936
25.0468 | -105.7961
25.0505 | -105.8001
25.0542 | -105.8047
25.0579 | -105.8095
25.0615 | -105.8142
25.0652 | -105.8201
25.0689 | -105.8253
25.0725 | -105.8291
25.0762 | -105.8334
25.0799 | -105.8386
25.0836 | -105.8442
25.0872 | -105.8523
25.0909 | -105.8597
25.0946 | -105.8665
25.0983 | -105.8736
25.1019 | -105.8805
25.1056 | -105.889
25.1093 | -105.898
25.1129 | -105.907
25.1166 | -105.9152
25.1203 | -105.9233
25.124 | -105.9335
25.1276 | -105.9439
25.1313 | -105.9519
25.135 | -105.9602
25.1387 | -105.9683
25.1423 | -105.9772
25.146 | -105.9839
25.1497 | -105.9902
25.1534 | -105.999
25.157 | -106.0092
25.1607 | -106.0182
25.1644 | -106.0244
25.168 | -106.0302
25.1717 | -106.0352
25.1754 | -106.0414
25.1791 | -106.0479
25.1827 | -106.0538
25.1864 | -106.0609
25.1901 | -106.0683
25.1938 | -106.0749
25.1974 | -106.0804
25.2011 | -106.0852
25.2048 | -106.0904
25.2084 | -106.0945
25.2121 | -106.0999
25.2158 | -106.1052
25.2195 | -106.1119
25.2231 | -106.1181
25.2268 | -106.1217
25.2305 | -106.1247
25.2342 | -106.129
25.2378 | -106.1338
25.2415 | -106.1386
25.2452 | -106.1423
25.2489 | -106.146
25.2525 | -106.15
25.2562 | -106.152
25.2599 | -106.1544
25.2635 | -106.1573
25.2672 | -106.1582
25.2709 | -106.1602
25.2746 | -106.1636
25.2782 | -106.166
25.2819 | -106.1676
25.2856 | -106.1683
25.2893 | -106.167
25.2929 | -106.1669
25.2966 | -106.1682
25.3003 | -106.1695
25.3039 | -106.1698
25.3076 | -106.1687
25.3113 | -106.1687
25.315 | -106.1675
25.3186 | -106.1657
25.3223 | -106.163
25.326 | -106.1605
25.3297 | -106.1577
25.3333 | -106.1534
25.337 | -106.1492
25.3407 | -106.1462
25.3444 | -106.141
25.348 | -106.136
25.3517 | -106.1336
25.3554 | -106.1303
25.359 | -106.1277
25.3627 | -106.1257
25.3664 | -106.1223
25.3701 | -106.119
25.3737 | -106.1169
25.3774 | -106.1139
25.3811 | -106.1106
25.3848 | -106.1053
25.3884 | -106.0989
25.3921 | -106.0937
25.3958 | -106.0886
25.3994 | -106.0837
25.4031 | -106.0809
25.4068 | -106.0766
25.4105 | -106.0724
25.4141 | -106.0681
25.4178 | -106.0631
25.4215 | -106.0608
25.4252 | -106.0594
25.4288 | -106.0586
25.4325 | -106.056
25.4362 | -106.0553
25.4399 | -106.0525
25.4435 | -106.0491
25.4472 | -106.045
25.4509 | -106.0429
25.4545 | -106.0404
25.4582 | -106.0376
25.4619 | -106.0355
25.4656 | -106.0323
25.4692 | -106.0289
25.4729 | -106.0254
25.4766 | -106.0224
25.4803 | -106.0168
25.4839 | -106.0112
25.4876 | -106.0061
25.4913 | -106.0029
25.4949 | -105.9991
25.4986 | -105.9945
25.5023 | -105.9903
25.506 | -105.9871
25.5096 | -105.9816
25.5133 | -105.9746
25.517 | -105.9686
25.5207 | -105.9637
25.5243 | -105.9595
25.528 | -105.9549
25.5317 | -105.9506
25.5354 | -105.9461
25.539 | -105.9418
25.5427 | -105.9378
25.5464 | -105.9334
25.55 | -105.9294
25.5537 | -105.9257
25.5574 | -105.924
25.5611 | -105.9226
25.5647 | -105.9219
25.5684 | -105.9206
25.5721 | -105.9179
25.5758 | -105.9167
25.5794 | -105.915
25.5831 | -105.9133
25.5868 | -105.9114
25.5905 | -105.9106
25.5941 | -105.9118
25.5978 | -105.9104
25.6015 | -105.909
25.6051 | -105.9096
25.6088 | -105.9116
25.6125 | -105.9127
25.6162 | -105.9137
25.6198 | -105.916
25.6235 | -105.9185
25.6272 | -105.9212
25.6309 | -105.9216
25.6345 | -105.9224
25.6382 | -105.924
25.6419 | -105.925
25.6455 | -105.9271
25.6492 | -105.9311
25.6529 | -105.9328
25.6566 | -105.9347
25.6602 | -105.9365
25.6639 | -105.9382
25.6676 | -105.9397
25.6713 | -105.9412
25.6749 | -105.9427
25.6786 | -105.944
25.6823 | -105.9447
25.6859 | -105.9455
25.6896 | -105.9469
25.6933 | -105.9477
25.697 | -105.9482
25.7006 | -105.9468
25.7043 | -105.9458
25.708 | -105.9436
25.7117 | -105.9395
25.7153 | -105.9374
25.719 | -105.9347
25.7227 | -105.9302
25.7264 | -105.9274
25.73 | -105.9254
25.7337 | -105.9216
25.7374 | -105.9177
25.741 | -105.914
25.7447 | -105.9111
25.7484 | -105.9109
25.7521 | -105.909
25.7557 | -105.906
25.7594 | -105.9034
25.7631 | -105.8981
25.7668 | -105.8946
25.7704 | -105.8923
25.7741 | -105.8917
25.7778 | -105.8924
25.7815 | -105.8901
25.7851 | -105.888
25.7888 | -105.8865
25.7925 | -105.8872
25.7961 | -105.8867
25.7998 | -105.8854
25.8035 | -105.8873
25.8072 | -105.8896
25.8108 | -105.8924
25.8145 | -105.895
25.8182 | -105.8986
25.8219 | -105.9027
25.8255 | -105.9067
25.8292 | -105.9126
25.8329 | -105.9194
25.8365 | -105.9249
25.8402 | -105.9305
25.8439 | -105.9387
25.8476 | -105.9467
25.8512 | -105.9573
25.8549 | -105.9672
25.8586 | -105.9772
25.8623 | -105.9876
25.8659 | -105.9965
25.8696 | -106.0067
25.8733 | -106.0199
25.877 | -106.0343
25.8806 | -106.0477
25.8843 | -106.0587
25.888 | -106.0704
25.8916 | -106.0826
25.8953 | -106.0945
25.899 | -106.107
25.9027 | -106.1194
25.9063 | -106.132
25.91 | -106.144
25.9137 | -106.1563
25.9174 | -106.1693
25.921 | -106.1812
25.9247 | -106.1948
25.9284 | -106.2068
25.932 | -106.219
25.9357 | -106.2332
25.9394 | -106.246
25.9431 | -106.2577
25.9467 | -106.2701
25.9504 | -106.2847
25.9541 | -106.2973
25.9578 | -106.3076
25.9614 | -106.3193
25.9651 | -106.3252
25.9688 | -106.3365
25.9725 | -106.3461
25.9761 | -106.3541
25.9798 | -106.3629
25.9835 | -106.3712
25.9871 | -106.3787
25.9908 | -106.3844
25.9945 | -106.3887
25.9982 | -106.3937
26.0018 | -106.399
26.0055 | -106.4033
26.0092 | -106.4079
26.0129 | -106.4133
26.0165 | -106.4173
26.0202 | -106.42
26.0239 | -106.422
26.0275 | -106.4246
26.0312 | -106.428
26.0349 | -106.4304
26.0386 | -106.4323
26.0422 | -106.4337
26.0459 | -106.4341
26.0496 | -106.4348
26.0533 | -106.4347
26.0569 | -106.4346
26.0606 | -106.4351
26.0643 | -106.4359
26.068 | -106.4347
26.0716 | -106.4348
26.0753 | -106.4366
26.079 | -106.4379
26.0826 | -106.4381
26.0863 | -106.4382
26.09 | -106.4377
26.0937 | -106.4377
26.0973 | -106.4396
26.101 | -106.4402
26.1047 | -106.4406
26.1084 | -106.4406
26.112 | -106.4425
26.1157 | -106.4416
26.1194 | -106.4408
26.123 | -106.4418
26.1267 | -106.4425
26.1304 | -106.4444
26.1341 | -106.4463
26.1377 | -106.447
26.1414 | -106.4464
26.1451 | -106.4452
26.1488 | -106.446
26.1524 | -106.4469
26.1561 | -106.4478
26.1598 | -106.4462
26.1635 | -106.4472
26.1671 | -106.4496
26.1708 | -106.4499
26.1745 | -106.4502
26.1781 | -106.4498
26.1818 | -106.4503
26.1855 | -106.4496
26.1892 | -106.4503
26.1928 | -106.4505
26.1965 | -106.4501
26.2002 | -106.4495
26.2039 | -106.4499
26.2075 | -106.4503
26.2112 | -106.4507
26.2149 | -106.4505
26.2185 | -106.4487
26.2222 | -106.4478
26.2259 | -106.4455
26.2296 | -106.4436
26.2332 | -106.4448
26.2369 | -106.4459
26.2406 | -106.4461
26.2443 | -106.4453
26.2479 | -106.4451
26.2516 | -106.4469
26.2553 | -106.4481
26.259 | -106.4469
26.2626 | -106.4468
26.2663 | -106.4479
26.27 | -106.4479
26.2736 | -106.4466
26.2773 | -106.4464
26.281 | -106.4448
26.2847 | -106.4423
26.2883 | -106.4407
26.292 | -106.4392
26.2957 | -106.4383
26.2994 | -106.4376
26.303 | -106.4363
26.3067 | -106.435
26.3104 | -106.4333
26.3141 | -106.4309
26.3177 | -106.4275
26.3214 | -106.4249
26.3251 | -106.4241
26.3287 | -106.4223
26.3324 | -106.4206
26.3361 | -106.419
26.3398 | -106.4174
26.3434 | -106.4172
26.3471 | -106.4159
26.3508 | -106.4125
26.3545 | -106.4097
26.3581 | -106.4069
26.3618 | -106.4057
26.3655 | -106.4064
26.3691 | -106.4064
26.3728 | -106.4054
26.3765 | -106.405
26.3802 | -106.4037
26.3838 | -106.4012
26.3875 | -106.3996
26.3912 | -106.3985
26.3949 | -106.3957
26.3985 | -106.3923
26.4022 | -106.3905
26.4059 | -106.3883
26.4095 | -106.3856
26.4132 | -106.3837
26.4169 | -106.384
26.4206 | -106.3842
26.4242 | -106.3853
26.4279 | -106.3853
26.4316 | -106.3848
26.4353 | -106.3865
26.4389 | -106.3911
26.4426 | -106.393
26.4463 | -106.3941
26.45 | -106.3959
26.4536 | -106.3977
26.4573 | -106.3992
26.461 | -106.4022
26.4646 | -106.4043
26.4683 | -106.4047
26.472 | -106.4067
26.4757 | -106.4112
26.4793 | -106.4161
26.483 | -106.4201
26.4867 | -106.426
26.4904 | -106.4328
26.494 | -106.4381
26.4977 | -106.4421
26.5014 | -106.4458
26.5051 | -106.4488
26.5087 | -106.4522
26.5124 | -106.4572
26.5161 | -106.4631
26.5197 | -106.4691
26.5234 | -106.473
26.5271 | -106.4763
26.5308 | -106.4795
26.5344 | -106.483
26.5381 | -106.4847
26.5418 | -106.4873
26.5455 | -106.4917
26.5491 | -106.4962
26.5528 | -106.5008
26.5565 | -106.5028
26.5601 | -106.5065
26.5638 | -106.5099
26.5675 | -106.5128
26.5712 | -106.5164
26.5748 | -106.5209
26.5785 | -106.5253
26.5822 | -106.5298
26.5859 | -106.5332
26.5895 | -106.5373
26.5932 | -106.541
26.5969 | -106.5437
26.6006 | -106.5479
26.6042 | -106.5518
26.6079 | -106.5548
26.6116 | -106.5563
26.6152 | -106.5577
26.6189 | -106.5592
26.6226 | -106.5602
26.6263 | -106.5615
26.6299 | -106.5624
26.6336 | -106.5651
26.6373 | -106.5656
26.641 | -106.5636
26.6446 | -106.5628
26.6483 | -106.5633
26.652 | -106.5639
26.6556 | -106.5668
26.6593 | -106.5691
26.663 | -106.5691
26.6667 | -106.5668
26.6703 | -106.5657
26.674 | -106.5644
26.6777 | -106.5624
26.6814 | -106.5612
26.685 | -106.5613
26.6887 | -106.5601
26.6924 | -106.5561
26.6961 | -106.5529
26.6997 | -106.549
26.7034 | -106.5462
26.7071 | -106.542
26.7107 | -106.5387
26.7144 | -106.5354
26.7181 | -106.5324
26.7218 | -106.5293
26.7254 | -106.5272
26.7291 | -106.524
26.7328 | -106.5218
26.7365 | -106.5177
26.7401 | -106.5126
26.7438 | -106.5079
26.7475 | -106.5022
26.7511 | -106.496
26.7548 | -106.4901
26.7585 | -106.4855
26.7622 | -106.4804
26.7658 | -106.4774
26.7695 | -106.4761
26.7732 | -106.4728
26.7769 | -106.4695
26.7805 | -106.4675
26.7842 | -106.4656
26.7879 | -106.4635
26.7916 | -106.4592
26.7952 | -106.4554
26.7989 | -106.4505
26.8026 | -106.4474
26.8062 | -106.4455
26.8099 | -106.4427
26.8136 | -106.4405
26.8173 | -106.439
26.8209 | -106.4373
26.8246 | -106.4353
26.8283 | -106.4319
26.832 | -106.4312
26.8356 | -106.4313
26.8393 | -106.4316
26.843 | -106.4329
26.8466 | -106.4356
26.8503 | -106.4373
26.854 | -106.4385
26.8577 | -106.4392
26.8613 | -106.4401
26.865 | -106.442
26.8687 | -106.4453
26.8724 | -106.4497
26.876 | -106.4545
26.8797 | -106.4585
26.8834 | -106.4636
26.8871 | -106.4678
26.8907 | -106.4719
26.8944 | -106.4763
26.8981 | -106.4809
26.9017 | -106.4864
26.9054 | -106.4917
26.9091 | -106.4962
26.9128 | -106.5019
26.9164 | -106.5085
26.9201 | -106.5134
26.9238 | -106.5172
26.9275 | -106.5212
26.9311 | -106.5257
26.9348 | -106.5307
26.9385 | -106.5352
26.9421 | -106.5409
26.9458 | -106.548
26.9495 | -106.5515
26.9532 | -106.5532
26.9568 | -106.5573
26.9605 | -106.5613
26.9642 | -106.5648
26.9679 | -106.5677
26.9715 | -106.5706
26.9752 | -106.5723
26.9789 | -106.5743
26.9826 | -106.5762
26.9862 | -106.5778
26.9899 | -106.5792
26.9936 | -106.5805
26.9972 | -106.5813
27.0009 | -106.5818
27.0046 | -106.5814
27.0083 | -106.5805
27.0119 | -106.579
27.0156 | -106.5779
27.0193 | -106.5764
27.023 | -106.5713
27.0266 | -106.5707
27.0303 | -106.5679
27.034 | -106.5636
27.0376 | -106.558
27.0413 | -106.5531
27.045 | -106.5495
27.0487 | -106.5461
27.0523 | -106.5408
27.056 | -106.5369
27.0597 | -106.5331
27.0634 | -106.5298
27.067 | -106.5245
27.0707 | -106.5181
27.0744 | -106.5147
27.0781 | -106.5114
27.0817 | -106.5082
27.0854 | -106.5049
27.0891 | -106.5026
27.0927 | -106.4977
27.0964 | -106.4931
27.1001 | -106.4879
27.1038 | -106.4846
27.1074 | -106.4816
27.1111 | -106.4788
27.1148 | -106.4756
27.1185 | -106.473
27.1221 | -106.4708
27.1258 | -106.4695
27.1295 | -106.4695
27.1332 | -106.4692
27.1368 | -106.4679
27.1405 | -106.4676
27.1442 | -106.4685
27.1478 | -106.4676
27.1515 | -106.4665
27.1552 | -106.4674
27.1589 | -106.4672
27.1625 | -106.4685
27.1662 | -106.4691
27.1699 | -106.4702
27.1736 | -106.4715
27.1772 | -106.4734
27.1809 | -106.4766
27.1846 | -106.4793
27.1882 | -106.4824
27.1919 | -106.484
27.1956 | -106.4842
27.1993 | -106.4864
27.2029 | -106.4899
27.2066 | -106.4936
27.2103 | -106.4984
27.214 | -106.5001
27.2176 | -106.5002
27.2213 | -106.4998
27.225 | -106.4992
27.2286 | -106.4989
27.2323 | -106.4977
27.236 | -106.4991
27.2397 | -106.5014
27.2433 | -106.5025
27.247 | -106.502
27.2507 | -106.5008
27.2544 | -106.4994
27.258 | -106.4983
27.2617 | -106.4977
27.2654 | -106.4964
27.2691 | -106.4954
27.2727 | -106.4933
27.2764 | -106.4899
27.2801 | -106.4873
27.2837 | -106.4844
27.2874 | -106.4826
27.2911 | -106.4805
27.2948 | -106.4771
27.2984 | -106.4727
27.3021 | -106.4682
27.3058 | -106.4644
27.3095 | -106.4593
27.3131 | -106.4545
27.3168 | -106.4514
27.3205 | -106.4482
27.3242 | -106.4421
27.3278 | -106.4344
27.3315 | -106.4284
27.3352 | -106.4233
27.3388 | -106.4155
27.3425 | -106.4073
27.3462 | -106.4006
27.3499 | -106.3937
27.3535 | -106.3869
27.3572 | -106.3803
27.3609 | -106.3734
27.3646 | -106.3667
27.3682 | -106.3607
27.3719 | -106.3539
27.3756 | -106.3476
27.3792 | -106.3419
27.3829 | -106.3348
27.3866 | -106.3285
27.3903 | -106.3223
27.3939 | -106.318
27.3976 | -106.3128
27.4013 | -106.3058
27.405 | -106.2982
27.4086 | -106.2913
27.4123 | -106.2857
27.416 | -106.2805
27.4197 | -106.2749
27.4233 | -106.2687
27.427 | -106.2623
27.4307 | -106.2551
27.4343 | -106.2479
27.438 | -106.2423
27.4417 | -106.2353
27.4454 | -106.229
27.449 | -106.221
27.4527 | -106.2148
27.4564 | -106.2103
27.4601 | -106.2054
27.4637 | -106.1984
27.4674 | -106.1916
27.4711 | -106.185
27.4747 | -106.1789
27.4784 | -106.1735
27.4821 | -106.168
27.4858 | -106.1639
27.4894 | -106.1593
27.4931 | -106.1544
27.4968 | -106.1506
27.5005 | -106.146
27.5041 | -106.1424
27.5078 | -106.1396
27.5115 | -106.1361
27.5152 | -106.1335
27.5188 | -106.132
27.5225 | -106.1301
27.5262 | -106.1279
27.5298 | -106.1254
27.5335 | -106.1238
27.5372 | -106.1212
27.5409 | -106.1198
27.5445 | -106.1192
27.5482 | -106.1171
27.5519 | -106.1207
27.5556 | -106.1213
27.5592 | -106.1222
27.5629 | -106.1247
27.5666 | -106.1279
27.5702 | -106.1306
27.5739 | -106.1336
27.5776 | -106.1364
27.5813 | -106.1402
27.5849 | -106.1446
27.5886 | -106.1501
27.5923 | -106.1542
27.596 | -106.1571
27.5996 | -106.1602
27.6033 | -106.1652
27.607 | -106.1719
27.6107 | -106.1791
27.6143 | -106.1859
27.618 | -106.1911
27.6217 | -106.1978
27.6253 | -106.203
27.629 | -106.2086
27.6327 | -106.2155
27.6364 | -106.2231
27.64 | -106.2303
27.6437 | -106.2364
27.6474 | -106.2421
27.6511 | -106.2466
27.6547 | -106.2515
27.6584 | -106.2564
27.6621 | -106.2607
27.6657 | -106.2664
27.6694 | -106.2716
27.6731 | -106.2759
27.6768 | -106.2794
27.6804 | -106.2819
27.6841 | -106.2838
27.6878 | -106.286
27.6915 | -106.2883
27.6951 | -106.2888
27.6988 | -106.2898
27.7025 | -106.2904
27.7062 | -106.2898
27.7098 | -106.2882
27.7135 | -106.2869
27.7172 | -106.2855
27.7208 | -106.2844
27.7245 | -106.2844
27.7282 | -106.2829
27.7319 | -106.2817
27.7355 | -106.2792
27.7392 | -106.2745
27.7429 | -106.27
27.7466 | -106.2657
27.7502 | -106.259
27.7539 | -106.2518
27.7576 | -106.2435
27.7612 | -106.2355
27.7649 | -106.227
27.7686 | -106.2181
27.7723 | -106.2075
27.7759 | -106.1976
27.7796 | -106.1894
27.7833 | -106.1781
27.787 | -106.1665
27.7906 | -106.1574
27.7943 | -106.1473
27.798 | -106.1347
27.8017 | -106.1212
27.8053 | -106.1088
27.809 | -106.0968
27.8127 | -106.0847
27.8163 | -106.0732
27.82 | -106.0598
27.8237 | -106.046
27.8274 | -106.0347
27.831 | -106.0238
27.8347 | -106.0124
27.8384 | -106.0022
27.8421 | -105.9915
27.8457 | -105.9803
27.8494 | -105.9687
27.8531 | -105.9583
27.8567 | -105.9457
27.8604 | -105.9309
27.8641 | -105.918
27.8678 | -105.9075
27.8714 | -105.8963
27.8751 | -105.8853
27.8788 | -105.8732
27.8825 | -105.8628
27.8861 | -105.8534
27.8898 | -105.8442
27.8935 | -105.8348
27.8972 | -105.8253
27.9008 | -105.8172
27.9045 | -105.8096
27.9082 | -105.8026
27.9118 | -105.7966
27.9155 | -105.7878
27.9192 | -105.7796
27.9229 | -105.7713
27.9265 | -105.7604
27.9302 | -105.7454
27.9339 | -105.7338
27.9376 | -105.727
27.9412 | -105.7196
27.9449 | -105.7107
27.9486 | -105.7031
27.9522 | -105.6952
27.9559 | -105.6883
27.9596 | -105.6799
27.9633 | -105.6718
27.9669 | -105.6651
27.9706 | -105.6575
27.9743 | -105.6505
27.978 | -105.6445
27.9816 | -105.6355
27.9853 | -105.6281
27.989 | -105.6233
27.9927 | -105.6178
27.9963 | -105.6097
28 | -105.6014
28.0037 | -105.5925