## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 285 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 290
Hood | 02-92B/P | 227
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1015.9 | hPa |  | PRC PRO 01
Humidity | 56.6 | %/Hr
Internal temperature of calibration bench | 24.21 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1022 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 23.7 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.96 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.99 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3085 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 10/09/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 12/09/2018
Technical Manager | Anthony HUE | Approval |  | 27/09/2018
CEO | Robert GLEMOT | Approval |  | 27/09/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4746 | 57.5106
0.012209 | -36.037 | 38.3443
0.018313 | -34.9777 | 28.2068
0.024417 | -34.5147 | 21.9657
0.030521 | -34.4032 | 18.2115
0.036626 | -34.258 | 15.1586
0.04273 | -34.1629 | 12.7843
0.048834 | -34.1176 | 11.2712
0.054938 | -34.0515 | 10.1488
0.061043 | -33.989 | 9.0686
0.067147 | -33.9788 | 8.2441
0.073251 | -33.9801 | 7.4838
0.079355 | -33.9991 | 6.8223
0.08546 | -34.0136 | 6.2281
0.091564 | -34.0084 | 5.7894
0.097668 | -33.9894 | 5.5445
0.10377 | -33.9668 | 5.2978
0.10988 | -33.9492 | 5.0208
0.11598 | -33.94 | 4.7271
0.12209 | -33.942 | 4.4422
0.12819 | -33.9522 | 4.1576
0.13429 | -33.9633 | 3.8685
0.1404 | -33.9665 | 3.5897
0.1465 | -33.9684 | 3.3317
0.15261 | -33.9655 | 3.1115
0.15871 | -33.9587 | 2.9499
0.16482 | -33.9503 | 2.8482
0.17092 | -33.9479 | 2.8105
0.17702 | -33.9456 | 2.7392
0.18313 | -33.9446 | 2.6616
0.18923 | -33.9442 | 2.5525
0.19534 | -33.9428 | 2.4253
0.20144 | -33.9415 | 2.2824
0.20754 | -33.9384 | 2.1614
0.21365 | -33.9374 | 2.0468
0.21975 | -33.9368 | 1.9491
0.22586 | -33.9378 | 1.862
0.23196 | -33.9402 | 1.7901
0.23807 | -33.9443 | 1.7279
0.24417 | -33.9484 | 1.6699
0.25027 | -33.9524 | 1.6081
0.25638 | -33.9549 | 1.5468
0.26248 | -33.9568 | 1.4805
0.26859 | -33.9577 | 1.4157
0.27469 | -33.9581 | 1.3447
0.2808 | -33.9591 | 1.2707
0.2869 | -33.9599 | 1.1996
0.293 | -33.9608 | 1.1324
0.29911 | -33.9616 | 1.0723
0.30521 | -33.9625 | 1.0131
0.31132 | -33.9631 | 0.97497
0.31742 | -33.9642 | 0.93255
0.32353 | -33.9644 | 0.89818
0.32963 | -33.9651 | 0.86989
0.33573 | -33.9661 | 0.85046
0.34184 | -33.9673 | 0.82994
0.34794 | -33.9694 | 0.80685
0.35405 | -33.9715 | 0.78346
0.36015 | -33.974 | 0.7587
0.36626 | -33.976 | 0.73168
0.37236 | -33.9781 | 0.69514
0.37846 | -33.9802 | 0.65077
0.38457 | -33.9811 | 0.59306
0.39067 | -33.9821 | 0.54612
0.39678 | -33.9828 | 0.49375
0.40288 | -33.9832 | 0.43912
0.40899 | -33.9831 | 0.3867
0.41509 | -33.9824 | 0.33534
0.42119 | -33.9824 | 0.29187
0.4273 | -33.9823 | 0.24534
0.4334 | -33.9822 | 0.19898
0.43951 | -33.9825 | 0.15549
0.44561 | -33.9834 | 0.11574
0.45172 | -33.9849 | 0.084384
0.45782 | -33.9857 | 0.04738
0.46392 | -33.9869 | 0.011358
0.47003 | -33.988 | -0.020835
0.47613 | -33.9888 | -0.054827
0.48224 | -33.9893 | -0.089544
0.48834 | -33.9883 | -0.1245
0.49445 | -33.988 | -0.15987
0.50055 | -33.9873 | -0.19208
0.50665 | -33.9865 | -0.22652
0.51276 | -33.9854 | -0.26082
0.51886 | -33.9843 | -0.2951
0.52497 | -33.983 | -0.31941
0.53107 | -33.9821 | -0.34908
0.53717 | -33.9813 | -0.37736
0.54328 | -33.9804 | -0.40022
0.54938 | -33.9801 | -0.42267
0.55549 | -33.9801 | -0.4436
0.56159 | -33.9812 | -0.45489
0.5677 | -33.9815 | -0.46926
0.5738 | -33.9825 | -0.4854
0.5799 | -33.9833 | -0.50011
0.58601 | -33.9844 | -0.50994
0.59211 | -33.9861 | -0.51598
0.59822 | -33.9875 | -0.52523
0.60432 | -33.989 | -0.53715
0.61043 | -33.9903 | -0.54483
0.61653 | -33.9916 | -0.555
0.62263 | -33.9931 | -0.56419
0.62874 | -33.9938 | -0.57466
0.63484 | -33.9946 | -0.5837
0.64095 | -33.9951 | -0.59369
0.64705 | -33.9955 | -0.60321
0.65316 | -33.9957 | -0.61553
0.65926 | -33.9957 | -0.63128
0.66536 | -33.9945 | -0.64859
0.67147 | -33.9941 | -0.66155
0.67757 | -33.9937 | -0.67569
0.68368 | -33.9929 | -0.69451
0.68978 | -33.9921 | -0.71151
0.69589 | -33.9911 | -0.73046
0.70199 | -33.9901 | -0.75323
0.70809 | -33.9892 | -0.77234
0.7142 | -33.9884 | -0.79279
0.7203 | -33.9873 | -0.8119
0.72641 | -33.9866 | -0.83447
0.73251 | -33.9859 | -0.85454
0.73862 | -33.9861 | -0.87649
0.74472 | -33.9857 | -0.8966
0.75082 | -33.9855 | -0.91886
0.75693 | -33.9858 | -0.94086
0.76303 | -33.9863 | -0.96301
0.76914 | -33.9882 | -0.98374
0.77524 | -33.9891 | -1.0035
0.78135 | -33.9904 | -1.0238
0.78745 | -33.9918 | -1.0439
0.79355 | -33.993 | -1.0638
0.79966 | -33.9948 | -1.0834
0.80576 | -33.9971 | -1.1032
0.81187 | -33.9988 | -1.1249
0.81797 | -34.0007 | -1.1464
0.82408 | -34.0025 | -1.1672
0.83018 | -34.0042 | -1.1879
0.83628 | -34.0062 | -1.2064
0.84239 | -34.0073 | -1.227
0.84849 | -34.0088 | -1.2488
0.8546 | -34.0103 | -1.2686
0.8607 | -34.0115 | -1.2894
0.86681 | -34.0126 | -1.3084
0.87291 | -34.0136 | -1.3272
0.87901 | -34.0141 | -1.3451
0.88512 | -34.0148 | -1.3632
0.89122 | -34.0155 | -1.3814
0.89733 | -34.0159 | -1.3969
0.90343 | -34.0163 | -1.4144
0.90953 | -34.0165 | -1.4286
0.91564 | -34.0165 | -1.4443
0.92174 | -34.0167 | -1.4592
0.92785 | -34.0167 | -1.4736
0.93395 | -34.0168 | -1.4872
0.94006 | -34.0169 | -1.4999
0.94616 | -34.0166 | -1.5116
0.95226 | -34.0167 | -1.524
0.95837 | -34.0168 | -1.5377
0.96447 | -34.0169 | -1.5504
0.97058 | -34.0167 | -1.5636
0.97668 | -34.0165 | -1.5767
0.98279 | -34.0162 | -1.5893
0.98889 | -34.0159 | -1.5997
0.99499 | -34.0159 | -1.613
1.0011 | -34.0157 | -1.6234
1.0072 | -34.0153 | -1.6349
1.0133 | -34.0149 | -1.6454
1.0194 | -34.0145 | -1.6582
1.0255 | -34.0141 | -1.6717
1.0316 | -34.0135 | -1.6834
1.0377 | -34.0131 | -1.697
1.0438 | -34.0125 | -1.7095
1.0499 | -34.0117 | -1.7247
1.056 | -34.0109 | -1.7389
1.0621 | -34.0103 | -1.753
1.0682 | -34.0098 | -1.7653
1.0743 | -34.0091 | -1.7807
1.0805 | -34.0086 | -1.795
1.0866 | -34.0076 | -1.81
1.0927 | -34.0067 | -1.8244
1.0988 | -34.0061 | -1.8392
1.1049 | -34.0052 | -1.8533
1.111 | -34.0044 | -1.8687
1.1171 | -34.0037 | -1.8832
1.1232 | -34.0032 | -1.8983
1.1293 | -34.0027 | -1.9123
1.1354 | -34.0024 | -1.9273
1.1415 | -34.002 | -1.9427
1.1476 | -34.0014 | -1.9582
1.1537 | -34.0008 | -1.9728
1.1598 | -34.0009 | -1.9876
1.1659 | -34.0006 | -2.0008
1.172 | -34.0004 | -2.0146
1.1781 | -34.0001 | -2.0285
1.1842 | -33.9999 | -2.0426
1.1903 | -34.0002 | -2.0534
1.1964 | -34.0001 | -2.0667
1.2025 | -34.0002 | -2.0805
1.2086 | -34.0005 | -2.0929
1.2147 | -34.0005 | -2.1053
1.2209 | -34.0008 | -2.1163
1.227 | -34.0015 | -2.1269
1.2331 | -34.0019 | -2.1387
1.2392 | -34.0022 | -2.1503
1.2453 | -34.0026 | -2.1611
1.2514 | -34.003 | -2.1708
1.2575 | -34.0036 | -2.1815
1.2636 | -34.0044 | -2.19
1.2697 | -34.005 | -2.199
1.2758 | -34.0056 | -2.2098
1.2819 | -34.0063 | -2.2194
1.288 | -34.007 | -2.2299
1.2941 | -34.0078 | -2.238
1.3002 | -34.0084 | -2.2457
1.3063 | -34.009 | -2.2545
1.3124 | -34.0096 | -2.2631
1.3185 | -34.0103 | -2.2718
1.3246 | -34.011 | -2.2805
1.3307 | -34.0118 | -2.2887
1.3368 | -34.0126 | -2.2971
1.3429 | -34.0133 | -2.3057
1.349 | -34.0137 | -2.3147
1.3551 | -34.0144 | -2.3236
1.3613 | -34.015 | -2.3315
1.3674 | -34.0156 | -2.3413
1.3735 | -34.016 | -2.3507
1.3796 | -34.0165 | -2.3594
1.3857 | -34.0168 | -2.3684
1.3918 | -34.0173 | -2.3783
1.3979 | -34.0179 | -2.3872
1.404 | -34.0184 | -2.3984
1.4101 | -34.0188 | -2.4086
1.4162 | -34.0192 | -2.4189
1.4223 | -34.0197 | -2.4301
1.4284 | -34.0201 | -2.4399
1.4345 | -34.0204 | -2.4512
1.4406 | -34.0207 | -2.4647
1.4467 | -34.0211 | -2.4761
1.4528 | -34.0216 | -2.4873
1.4589 | -34.0221 | -2.499
1.465 | -34.0223 | -2.5097
1.4711 | -34.0228 | -2.521
1.4772 | -34.0233 | -2.5343
1.4833 | -34.0238 | -2.5471
1.4894 | -34.0243 | -2.5592
1.4955 | -34.0249 | -2.5711
1.5016 | -34.0254 | -2.5835
1.5078 | -34.0258 | -2.5974
1.5139 | -34.0262 | -2.6104
1.52 | -34.0266 | -2.6229
1.5261 | -34.027 | -2.6355
1.5322 | -34.0273 | -2.6495
1.5383 | -34.0277 | -2.662
1.5444 | -34.0282 | -2.6746
1.5505 | -34.0287 | -2.688
1.5566 | -34.0292 | -2.701
1.5627 | -34.0297 | -2.7148
1.5688 | -34.0303 | -2.7283
1.5749 | -34.0307 | -2.741
1.581 | -34.031 | -2.7544
1.5871 | -34.0314 | -2.7681
1.5932 | -34.0317 | -2.781
1.5993 | -34.0321 | -2.7952
1.6054 | -34.0325 | -2.8085
1.6115 | -34.0329 | -2.8226
1.6176 | -34.0332 | -2.8352
1.6237 | -34.0334 | -2.8489
1.6298 | -34.0337 | -2.8615
1.6359 | -34.0341 | -2.8743
1.642 | -34.0345 | -2.8863
1.6482 | -34.0347 | -2.8994
1.6543 | -34.035 | -2.9131
1.6604 | -34.0353 | -2.9251
1.6665 | -34.0354 | -2.9377
1.6726 | -34.0356 | -2.9502
1.6787 | -34.0359 | -2.9629
1.6848 | -34.0358 | -2.9756
1.6909 | -34.0361 | -2.9884
1.697 | -34.0363 | -3.0006
1.7031 | -34.0363 | -3.0133
1.7092 | -34.0363 | -3.0256
1.7153 | -34.0365 | -3.0376
1.7214 | -34.0362 | -3.0498
1.7275 | -34.0362 | -3.0625
1.7336 | -34.0361 | -3.0755
1.7397 | -34.0362 | -3.0878
1.7458 | -34.0362 | -3.1001
1.7519 | -34.0362 | -3.1123
1.758 | -34.0362 | -3.1243
1.7641 | -34.0364 | -3.1359
1.7702 | -34.0363 | -3.1489
1.7763 | -34.0363 | -3.1605
1.7824 | -34.0362 | -3.173
1.7885 | -34.0362 | -3.1854
1.7947 | -34.0363 | -3.1975
1.8008 | -34.0363 | -3.2088
1.8069 | -34.0363 | -3.2204
1.813 | -34.0363 | -3.2326
1.8191 | -34.0363 | -3.2453
1.8252 | -34.0364 | -3.258
1.8313 | -34.0365 | -3.2702
1.8374 | -34.0367 | -3.282
1.8435 | -34.0367 | -3.2939
1.8496 | -34.0368 | -3.307
1.8557 | -34.0369 | -3.3197
1.8618 | -34.0372 | -3.3318
1.8679 | -34.0373 | -3.3437
1.874 | -34.0374 | -3.3556
1.8801 | -34.0376 | -3.3679
1.8862 | -34.0378 | -3.3799
1.8923 | -34.038 | -3.3919
1.8984 | -34.0381 | -3.4039
1.9045 | -34.0382 | -3.4153
1.9106 | -34.0382 | -3.4274
1.9167 | -34.0382 | -3.4395
1.9228 | -34.0384 | -3.4512
1.9289 | -34.0385 | -3.4629
1.9351 | -34.0387 | -3.4736
1.9412 | -34.0389 | -3.4847
1.9473 | -34.0392 | -3.4966
1.9534 | -34.0394 | -3.5079
1.9595 | -34.0395 | -3.5196
1.9656 | -34.0398 | -3.5317
1.9717 | -34.04 | -3.5432
1.9778 | -34.0404 | -3.5545
1.9839 | -34.0406 | -3.5658
1.99 | -34.0408 | -3.5766
1.9961 | -34.0409 | -3.5886
2.0022 | -34.0411 | -3.5999
2.0083 | -34.0412 | -3.6109
2.0144 | -34.0414 | -3.6222
2.0205 | -34.0417 | -3.6338
2.0266 | -34.042 | -3.645
2.0327 | -34.0423 | -3.6564
2.0388 | -34.0427 | -3.6678
2.0449 | -34.0429 | -3.679
2.051 | -34.0431 | -3.6902
2.0571 | -34.0435 | -3.7014
2.0632 | -34.0436 | -3.7128
2.0693 | -34.0438 | -3.7248
2.0754 | -34.0441 | -3.736
2.0816 | -34.0444 | -3.748
2.0877 | -34.0445 | -3.7594
2.0938 | -34.0449 | -3.7713
2.0999 | -34.0451 | -3.7829
2.106 | -34.0452 | -3.7944
2.1121 | -34.0454 | -3.8063
2.1182 | -34.0457 | -3.8169
2.1243 | -34.0459 | -3.8275
2.1304 | -34.0461 | -3.8385
2.1365 | -34.0464 | -3.8501
2.1426 | -34.0466 | -3.862
2.1487 | -34.0468 | -3.8731
2.1548 | -34.0472 | -3.8848
2.1609 | -34.0477 | -3.8961
2.167 | -34.048 | -3.9072
2.1731 | -34.0483 | -3.9191
2.1792 | -34.0485 | -3.93
2.1853 | -34.0489 | -3.9419
2.1914 | -34.0492 | -3.9531
2.1975 | -34.0495 | -3.9644
2.2036 | -34.0497 | -3.9764
2.2097 | -34.0499 | -3.9877
2.2158 | -34.0503 | -3.9982
2.222 | -34.0505 | -4.0085
2.2281 | -34.0507 | -4.0198
2.2342 | -34.051 | -4.031
2.2403 | -34.0512 | -4.0415
2.2464 | -34.0515 | -4.0522
2.2525 | -34.0517 | -4.0635
2.2586 | -34.052 | -4.0739
2.2647 | -34.0523 | -4.0842
2.2708 | -34.0526 | -4.0953
2.2769 | -34.0529 | -4.1058
2.283 | -34.053 | -4.1169
2.2891 | -34.0532 | -4.1276
2.2952 | -34.0534 | -4.1381
2.3013 | -34.0536 | -4.1488
2.3074 | -34.0538 | -4.1591
2.3135 | -34.054 | -4.1704
2.3196 | -34.0542 | -4.1804
2.3257 | -34.0544 | -4.1909
2.3318 | -34.0545 | -4.2018
2.3379 | -34.0547 | -4.2124
2.344 | -34.0548 | -4.2237
2.3501 | -34.055 | -4.2343
2.3562 | -34.0551 | -4.2446
2.3623 | -34.0552 | -4.2548
2.3685 | -34.0553 | -4.2652
2.3746 | -34.0554 | -4.2763
2.3807 | -34.0556 | -4.2866
2.3868 | -34.0558 | -4.2972
2.3929 | -34.0559 | -4.3078
2.399 | -34.056 | -4.318
2.4051 | -34.0563 | -4.3284
2.4112 | -34.0564 | -4.3389
2.4173 | -34.0566 | -4.3494
2.4234 | -34.0566 | -4.3601
2.4295 | -34.0568 | -4.3708
2.4356 | -34.0569 | -4.3815
2.4417 | -34.057 | -4.3925
2.4478 | -34.0571 | -4.4033
2.4539 | -34.0573 | -4.4134
2.46 | -34.0574 | -4.4247
2.4661 | -34.0575 | -4.4357
2.4722 | -34.0577 | -4.4458
2.4783 | -34.0577 | -4.456
2.4844 | -34.0578 | -4.467
2.4905 | -34.0579 | -4.4778
2.4966 | -34.058 | -4.4884
2.5027 | -34.0581 | -4.4987
2.5089 | -34.0582 | -4.5095
2.515 | -34.0583 | -4.5201
2.5211 | -34.0583 | -4.5308
2.5272 | -34.0584 | -4.5416
2.5333 | -34.0585 | -4.5524
2.5394 | -34.0587 | -4.5631
2.5455 | -34.0588 | -4.5743
2.5516 | -34.059 | -4.5845
2.5577 | -34.0591 | -4.5959
2.5638 | -34.0593 | -4.6076
2.5699 | -34.0593 | -4.6189
2.576 | -34.0594 | -4.6304
2.5821 | -34.0595 | -4.6415
2.5882 | -34.0596 | -4.6526
2.5943 | -34.0597 | -4.6633
2.6004 | -34.0599 | -4.6742
2.6065 | -34.06 | -4.6844
2.6126 | -34.0601 | -4.6948
2.6187 | -34.0602 | -4.7052
2.6248 | -34.0604 | -4.7159
2.6309 | -34.0605 | -4.7273
2.637 | -34.0608 | -4.7389
2.6431 | -34.0609 | -4.7493
2.6492 | -34.061 | -4.7606
2.6554 | -34.0612 | -4.7713
2.6615 | -34.0613 | -4.7823
2.6676 | -34.0616 | -4.7936
2.6737 | -34.0619 | -4.8047
2.6798 | -34.0622 | -4.8158
2.6859 | -34.0623 | -4.827
2.692 | -34.0625 | -4.8378
2.6981 | -34.0627 | -4.8488
2.7042 | -34.0629 | -4.8599
2.7103 | -34.0631 | -4.8704
2.7164 | -34.0634 | -4.8813
2.7225 | -34.0636 | -4.8926
2.7286 | -34.0639 | -4.9036
2.7347 | -34.0641 | -4.9141
2.7408 | -34.0644 | -4.9241
2.7469 | -34.0646 | -4.9353
2.753 | -34.0649 | -4.9458
2.7591 | -34.0652 | -4.957
2.7652 | -34.0654 | -4.9685
2.7713 | -34.0656 | -4.9792
2.7774 | -34.0658 | -4.99
2.7835 | -34.066 | -5.001
2.7896 | -34.0662 | -5.0114
2.7958 | -34.0665 | -5.0221
2.8019 | -34.0668 | -5.0324
2.808 | -34.067 | -5.0432
2.8141 | -34.0674 | -5.0543
2.8202 | -34.0677 | -5.0651
2.8263 | -34.068 | -5.0759
2.8324 | -34.0683 | -5.0873
2.8385 | -34.0687 | -5.0981
2.8446 | -34.0691 | -5.1087
2.8507 | -34.0694 | -5.1195
2.8568 | -34.0698 | -5.1299
2.8629 | -34.07 | -5.1404
2.869 | -34.0704 | -5.1514
2.8751 | -34.0707 | -5.1621
2.8812 | -34.0711 | -5.1733
2.8873 | -34.0714 | -5.1841
2.8934 | -34.0717 | -5.1951
2.8995 | -34.072 | -5.2061
2.9056 | -34.0723 | -5.2172
2.9117 | -34.0726 | -5.2281
2.9178 | -34.073 | -5.2388
2.9239 | -34.0734 | -5.2503
2.93 | -34.0738 | -5.2617
2.9361 | -34.0741 | -5.2724
2.9423 | -34.0744 | -5.2829
2.9484 | -34.0748 | -5.2939
2.9545 | -34.0752 | -5.3048
2.9606 | -34.0756 | -5.3162
2.9667 | -34.0759 | -5.3272
2.9728 | -34.0763 | -5.3384
2.9789 | -34.0767 | -5.3494
2.985 | -34.0771 | -5.3606
2.9911 | -34.0775 | -5.3717
2.9972 | -34.0779 | -5.3828
3.0033 | -34.0783 | -5.3942
3.0094 | -34.0787 | -5.4052
3.0155 | -34.0791 | -5.416
3.0216 | -34.0795 | -5.4271
3.0277 | -34.0798 | -5.4387
3.0338 | -34.0802 | -5.45
3.0399 | -34.0806 | -5.4612
3.046 | -34.0809 | -5.4727
3.0521 | -34.0812 | -5.4836
3.0582 | -34.0815 | -5.4955
3.0643 | -34.0818 | -5.5069
3.0704 | -34.0821 | -5.5186
3.0765 | -34.0824 | -5.5295
3.0827 | -34.0827 | -5.5406
3.0888 | -34.083 | -5.5518
3.0949 | -34.0834 | -5.5629
3.101 | -34.0837 | -5.5743
3.1071 | -34.084 | -5.5859
3.1132 | -34.0844 | -5.5976
3.1193 | -34.0847 | -5.6089
3.1254 | -34.0851 | -5.6201
3.1315 | -34.0855 | -5.6319
3.1376 | -34.0857 | -5.6432
3.1437 | -34.0861 | -5.6546
3.1498 | -34.0864 | -5.6666
3.1559 | -34.0867 | -5.6779
3.162 | -34.0869 | -5.6898
3.1681 | -34.0872 | -5.7016
3.1742 | -34.0875 | -5.713
3.1803 | -34.0878 | -5.7246
3.1864 | -34.0882 | -5.7357
3.1925 | -34.0884 | -5.7471
3.1986 | -34.0887 | -5.7589
3.2047 | -34.0889 | -5.7708
3.2108 | -34.0892 | -5.7817
3.2169 | -34.0895 | -5.7933
3.223 | -34.0897 | -5.8048
3.2292 | -34.0899 | -5.8165
3.2353 | -34.0901 | -5.8272
3.2414 | -34.0903 | -5.8384
3.2475 | -34.0905 | -5.8497
3.2536 | -34.0908 | -5.861
3.2597 | -34.091 | -5.8718
3.2658 | -34.0912 | -5.8826
3.2719 | -34.0914 | -5.8935
3.278 | -34.0917 | -5.905
3.2841 | -34.0919 | -5.9159
3.2902 | -34.0921 | -5.9272
3.2963 | -34.0924 | -5.9385
3.3024 | -34.0924 | -5.9495
3.3085 | -34.0927 | -5.9609
3.3146 | -34.0929 | -5.9719
3.3207 | -34.0931 | -5.9833
3.3268 | -34.0933 | -5.9946
3.3329 | -34.0935 | -6.0051
3.339 | -34.0936 | -6.0158
3.3451 | -34.0938 | -6.0269
3.3512 | -34.094 | -6.0374
3.3573 | -34.0942 | -6.0485
3.3634 | -34.0943 | -6.0595
3.3696 | -34.0945 | -6.0707
3.3757 | -34.0946 | -6.0814
3.3818 | -34.0949 | -6.0921
3.3879 | -34.095 | -6.1031
3.394 | -34.0952 | -6.1144
3.4001 | -34.0953 | -6.1251
3.4062 | -34.0954 | -6.1358
3.4123 | -34.0956 | -6.1463
3.4184 | -34.0958 | -6.1575
3.4245 | -34.096 | -6.1681
3.4306 | -34.096 | -6.1787
3.4367 | -34.0962 | -6.1891
3.4428 | -34.0964 | -6.1994
3.4489 | -34.0965 | -6.2102
3.455 | -34.0966 | -6.221
3.4611 | -34.0968 | -6.2321
3.4672 | -34.0969 | -6.2431
3.4733 | -34.0971 | -6.2541
3.4794 | -34.0973 | -6.2648
3.4855 | -34.0974 | -6.2751
3.4916 | -34.0976 | -6.2854
3.4977 | -34.0977 | -6.2966
3.5038 | -34.0979 | -6.3073
3.5099 | -34.098 | -6.3175
3.5161 | -34.0981 | -6.3278
3.5222 | -34.0982 | -6.3381
3.5283 | -34.0983 | -6.3488
3.5344 | -34.0984 | -6.3596
3.5405 | -34.0986 | -6.37
3.5466 | -34.0988 | -6.3809
3.5527 | -34.0989 | -6.3917
3.5588 | -34.099 | -6.4022
3.5649 | -34.099 | -6.4128
3.571 | -34.0991 | -6.4234
3.5771 | -34.0993 | -6.4344
3.5832 | -34.0994 | -6.4442
3.5893 | -34.0996 | -6.4548
3.5954 | -34.0998 | -6.4655
3.6015 | -34.0999 | -6.4764
3.6076 | -34.1001 | -6.4868
3.6137 | -34.1003 | -6.4968
3.6198 | -34.1005 | -6.507
3.6259 | -34.1007 | -6.5179
3.632 | -34.1008 | -6.5285
3.6381 | -34.101 | -6.5392
3.6442 | -34.1011 | -6.5497
3.6503 | -34.1014 | -6.5598
3.6565 | -34.1016 | -6.5706
3.6626 | -34.1017 | -6.5811
3.6687 | -34.1019 | -6.5917
3.6748 | -34.1022 | -6.602
3.6809 | -34.1024 | -6.6127
3.687 | -34.1026 | -6.6236
3.6931 | -34.1028 | -6.6337
3.6992 | -34.103 | -6.6444
3.7053 | -34.1032 | -6.6549
3.7114 | -34.1033 | -6.6655
3.7175 | -34.1036 | -6.6762
3.7236 | -34.1038 | -6.6862
3.7297 | -34.104 | -6.6969
3.7358 | -34.1042 | -6.7078
3.7419 | -34.1044 | -6.7186
3.748 | -34.1046 | -6.7291
3.7541 | -34.1048 | -6.7396
3.7602 | -34.1051 | -6.7501
3.7663 | -34.1053 | -6.7606
3.7724 | -34.1055 | -6.7712
3.7785 | -34.1058 | -6.7816
3.7846 | -34.106 | -6.792
3.7907 | -34.1063 | -6.8029
3.7969 | -34.1065 | -6.8136
3.803 | -34.1068 | -6.8237
3.8091 | -34.107 | -6.8342
3.8152 | -34.1073 | -6.8448
3.8213 | -34.1075 | -6.8556
3.8274 | -34.1077 | -6.8659
3.8335 | -34.1081 | -6.8766
3.8396 | -34.1084 | -6.8872
3.8457 | -34.1086 | -6.8979
3.8518 | -34.109 | -6.9084
3.8579 | -34.1093 | -6.9189
3.864 | -34.1096 | -6.9297
3.8701 | -34.1099 | -6.9406
3.8762 | -34.1101 | -6.951
3.8823 | -34.1103 | -6.9618
3.8884 | -34.1106 | -6.9724
3.8945 | -34.111 | -6.9833
3.9006 | -34.1113 | -6.9942
3.9067 | -34.1115 | -7.0053
3.9128 | -34.1118 | -7.016
3.9189 | -34.1121 | -7.0265
3.925 | -34.1123 | -7.0372
3.9311 | -34.1126 | -7.0474
3.9372 | -34.1129 | -7.0583
3.9434 | -34.1131 | -7.0689
3.9495 | -34.1135 | -7.0796
3.9556 | -34.1138 | -7.0901
3.9617 | -34.1141 | -7.1003
3.9678 | -34.1144 | -7.1107
3.9739 | -34.1147 | -7.1212
3.98 | -34.1149 | -7.1319
3.9861 | -34.1152 | -7.1426
3.9922 | -34.1154 | -7.1532
3.9983 | -34.1157 | -7.1634
4.0044 | -34.116 | -7.1736
4.0105 | -34.1162 | -7.1842
4.0166 | -34.1166 | -7.1949
4.0227 | -34.1168 | -7.2055
4.0288 | -34.1171 | -7.2165
4.0349 | -34.1173 | -7.2269
4.041 | -34.1175 | -7.2374
4.0471 | -34.1177 | -7.2481
4.0532 | -34.118 | -7.2583
4.0593 | -34.1183 | -7.2691
4.0654 | -34.1186 | -7.2794
4.0715 | -34.1189 | -7.2905
4.0776 | -34.1192 | -7.301
4.0838 | -34.1195 | -7.312
4.0899 | -34.1197 | -7.3223
4.096 | -34.12 | -7.3329
4.1021 | -34.1202 | -7.3432
4.1082 | -34.1205 | -7.3535
4.1143 | -34.1207 | -7.3637
4.1204 | -34.121 | -7.374
4.1265 | -34.1212 | -7.3842
4.1326 | -34.1215 | -7.3945
4.1387 | -34.1218 | -7.4049
4.1448 | -34.122 | -7.4155
4.1509 | -34.1222 | -7.426
4.157 | -34.1225 | -7.4369
4.1631 | -34.1227 | -7.4476
4.1692 | -34.123 | -7.458
4.1753 | -34.1233 | -7.4684
4.1814 | -34.1234 | -7.4786
4.1875 | -34.1237 | -7.4893
4.1936 | -34.124 | -7.4999
4.1997 | -34.1242 | -7.51
4.2058 | -34.1244 | -7.5207
4.2119 | -34.1246 | -7.531
4.218 | -34.1248 | -7.5412
4.2241 | -34.125 | -7.5514
4.2303 | -34.1253 | -7.5617
4.2364 | -34.1255 | -7.5721
4.2425 | -34.1257 | -7.5826
4.2486 | -34.1259 | -7.5931
4.2547 | -34.126 | -7.6033
4.2608 | -34.1262 | -7.6138
4.2669 | -34.1264 | -7.6239
4.273 | -34.1267 | -7.6344
4.2791 | -34.1269 | -7.6448
4.2852 | -34.1271 | -7.6554
4.2913 | -34.1272 | -7.6655
4.2974 | -34.1274 | -7.6761
4.3035 | -34.1277 | -7.6868
4.3096 | -34.1278 | -7.6974
4.3157 | -34.1281 | -7.7076
4.3218 | -34.1283 | -7.7178
4.3279 | -34.1285 | -7.7282
4.334 | -34.1286 | -7.7384
4.3401 | -34.1288 | -7.7484
4.3462 | -34.1291 | -7.7588
4.3523 | -34.1293 | -7.7687
4.3584 | -34.1295 | -7.7787
4.3645 | -34.1297 | -7.7889
4.3707 | -34.1299 | -7.7989
4.3768 | -34.1301 | -7.8092
4.3829 | -34.1303 | -7.8198
4.389 | -34.1305 | -7.83
4.3951 | -34.1307 | -7.8401
4.4012 | -34.1308 | -7.8507
4.4073 | -34.131 | -7.8608
4.4134 | -34.1312 | -7.871
4.4195 | -34.1315 | -7.8812
4.4256 | -34.1317 | -7.8912
4.4317 | -34.1319 | -7.9014
4.4378 | -34.1321 | -7.9113
4.4439 | -34.1323 | -7.9212
4.45 | -34.1324 | -7.9313
4.4561 | -34.1325 | -7.9415
4.4622 | -34.1327 | -7.9518
4.4683 | -34.1329 | -7.9619
4.4744 | -34.1331 | -7.9725
4.4805 | -34.1333 | -7.9822
4.4866 | -34.1335 | -7.9924
4.4927 | -34.1337 | -8.0028
4.4988 | -34.1339 | -8.0129
4.5049 | -34.1341 | -8.0227
4.511 | -34.1344 | -8.0326
4.5172 | -34.1346 | -8.0428
4.5233 | -34.1347 | -8.0533
4.5294 | -34.135 | -8.0633
4.5355 | -34.1352 | -8.0735
4.5416 | -34.1354 | -8.0836
4.5477 | -34.1356 | -8.0938
4.5538 | -34.1358 | -8.1035
4.5599 | -34.136 | -8.1136
4.566 | -34.1362 | -8.1238
4.5721 | -34.1364 | -8.1339
4.5782 | -34.1367 | -8.1443
4.5843 | -34.1369 | -8.1545
4.5904 | -34.1372 | -8.1648
4.5965 | -34.1374 | -8.1747
4.6026 | -34.1376 | -8.1853
4.6087 | -34.1378 | -8.1953
4.6148 | -34.138 | -8.2058
4.6209 | -34.1382 | -8.2159
4.627 | -34.1385 | -8.2262
4.6331 | -34.1386 | -8.2363
4.6392 | -34.1388 | -8.2465
4.6453 | -34.1391 | -8.2567
4.6514 | -34.1393 | -8.2668
4.6576 | -34.1395 | -8.2774
4.6637 | -34.1397 | -8.2872
4.6698 | -34.1399 | -8.2972
4.6759 | -34.1401 | -8.3078
4.682 | -34.1403 | -8.3179
4.6881 | -34.1404 | -8.328
4.6942 | -34.1407 | -8.3381
4.7003 | -34.1409 | -8.3483
4.7064 | -34.1411 | -8.3583
4.7125 | -34.1414 | -8.3688
4.7186 | -34.1415 | -8.3788
4.7247 | -34.1418 | -8.3891
4.7308 | -34.142 | -8.3992
4.7369 | -34.1422 | -8.4095
4.743 | -34.1425 | -8.4201
4.7491 | -34.1427 | -8.4306
4.7552 | -34.1429 | -8.4409
4.7613 | -34.1431 | -8.4513
4.7674 | -34.1433 | -8.4613
4.7735 | -34.1435 | -8.472
4.7796 | -34.1438 | -8.4825
4.7857 | -34.144 | -8.4928
4.7918 | -34.1442 | -8.5029
4.7979 | -34.1445 | -8.513
4.8041 | -34.1447 | -8.5233
4.8102 | -34.145 | -8.5332
4.8163 | -34.1452 | -8.5434
4.8224 | -34.1454 | -8.5536
4.8285 | -34.1457 | -8.5635
4.8346 | -34.146 | -8.5736
4.8407 | -34.1461 | -8.5839
4.8468 | -34.1464 | -8.5945
4.8529 | -34.1467 | -8.605
4.859 | -34.1469 | -8.6156
4.8651 | -34.1471 | -8.6256
4.8712 | -34.1474 | -8.6357
4.8773 | -34.1475 | -8.6461
4.8834 | -34.1478 | -8.6563
4.8895 | -34.148 | -8.6664
4.8956 | -34.1482 | -8.6766
4.9017 | -34.1484 | -8.6868
4.9078 | -34.1487 | -8.6974
4.9139 | -34.1489 | -8.7072
4.92 | -34.1492 | -8.7172
4.9261 | -34.1495 | -8.7271
4.9322 | -34.1497 | -8.7374
4.9383 | -34.15 | -8.7476
4.9445 | -34.1503 | -8.7579
4.9506 | -34.1505 | -8.7679
4.9567 | -34.1507 | -8.778
4.9628 | -34.1509 | -8.7882
4.9689 | -34.1512 | -8.7989
4.975 | -34.1514 | -8.8095
4.9811 | -34.1517 | -8.8194
4.9872 | -34.152 | -8.8295
4.9933 | -34.1523 | -8.8401
4.9994 | -34.1526 | -8.8508
5.0055 | -34.1528 | -8.8611
5.0116 | -34.1531 | -8.8718
5.0177 | -34.1534 | -8.8823
5.0238 | -34.1537 | -8.8929
5.0299 | -34.1541 | -8.9036
5.036 | -34.1543 | -8.9138
5.0421 | -34.1545 | -8.924
5.0482 | -34.1548 | -8.9343
5.0543 | -34.1551 | -8.9444
5.0604 | -34.1554 | -8.9547
5.0665 | -34.1557 | -8.9647
5.0726 | -34.156 | -8.9754
5.0787 | -34.1563 | -8.9859
5.0848 | -34.1565 | -8.9963
5.091 | -34.1568 | -9.0065
5.0971 | -34.1571 | -9.0164
5.1032 | -34.1574 | -9.0271
5.1093 | -34.1577 | -9.0374
5.1154 | -34.1579 | -9.0475
5.1215 | -34.1582 | -9.0576
5.1276 | -34.1584 | -9.0678
5.1337 | -34.1588 | -9.0778
5.1398 | -34.159 | -9.0879
5.1459 | -34.1593 | -9.0985
5.152 | -34.1596 | -9.109
5.1581 | -34.1599 | -9.1194
5.1642 | -34.1602 | -9.1298
5.1703 | -34.1605 | -9.14
5.1764 | -34.1608 | -9.1502
5.1825 | -34.1611 | -9.1603
5.1886 | -34.1614 | -9.1702
5.1947 | -34.1616 | -9.1805
5.2008 | -34.1619 | -9.1909
5.2069 | -34.1622 | -9.2014
5.213 | -34.1625 | -9.2115
5.2191 | -34.1628 | -9.2219
5.2252 | -34.1631 | -9.2322
5.2314 | -34.1633 | -9.2428
5.2375 | -34.1636 | -9.2528
5.2436 | -34.1639 | -9.2631
5.2497 | -34.1642 | -9.2731
5.2558 | -34.1645 | -9.2838
5.2619 | -34.1648 | -9.2944
5.268 | -34.1651 | -9.3049
5.2741 | -34.1654 | -9.3153
5.2802 | -34.1657 | -9.3259
5.2863 | -34.1661 | -9.3362
5.2924 | -34.1664 | -9.3467
5.2985 | -34.1667 | -9.3571
5.3046 | -34.167 | -9.3675
5.3107 | -34.1673 | -9.378
5.3168 | -34.1676 | -9.3884
5.3229 | -34.1678 | -9.3987
5.329 | -34.1681 | -9.4091
5.3351 | -34.1684 | -9.4191
5.3412 | -34.1687 | -9.4294
5.3473 | -34.169 | -9.4398
5.3534 | -34.1693 | -9.4502
5.3595 | -34.1696 | -9.4606
5.3656 | -34.17 | -9.4709
5.3717 | -34.1702 | -9.4815
5.3779 | -34.1705 | -9.4922
5.384 | -34.1708 | -9.5025
5.3901 | -34.1712 | -9.5132
5.3962 | -34.1714 | -9.5233
5.4023 | -34.1718 | -9.5339
5.4084 | -34.1721 | -9.5442
5.4145 | -34.1724 | -9.5543
5.4206 | -34.1727 | -9.5647
5.4267 | -34.173 | -9.5754
5.4328 | -34.1733 | -9.5857
5.4389 | -34.1736 | -9.5958
5.445 | -34.1739 | -9.6065
5.4511 | -34.1742 | -9.6167
5.4572 | -34.1745 | -9.6272
5.4633 | -34.1748 | -9.6375
5.4694 | -34.1751 | -9.6479
5.4755 | -34.1754 | -9.658
5.4816 | -34.1757 | -9.6684
5.4877 | -34.176 | -9.6786
5.4938 | -34.1763 | -9.6887
5.4999 | -34.1766 | -9.6989
5.506 | -34.177 | -9.7092
5.5121 | -34.1772 | -9.7195
5.5183 | -34.1775 | -9.7295
5.5244 | -34.1778 | -9.7401
5.5305 | -34.1781 | -9.7503
5.5366 | -34.1784 | -9.7604
5.5427 | -34.1787 | -9.7708
5.5488 | -34.179 | -9.7812
5.5549 | -34.1793 | -9.7915
5.561 | -34.1796 | -9.8021
5.5671 | -34.1799 | -9.8123
5.5732 | -34.1802 | -9.8228
5.5793 | -34.1805 | -9.8334
5.5854 | -34.1808 | -9.8437
5.5915 | -34.1811 | -9.8543
5.5976 | -34.1814 | -9.8647
5.6037 | -34.1817 | -9.8751
5.6098 | -34.182 | -9.8854
5.6159 | -34.1823 | -9.8955
5.622 | -34.1826 | -9.9059
5.6281 | -34.1829 | -9.9163
5.6342 | -34.1832 | -9.927
5.6403 | -34.1835 | -9.9374
5.6464 | -34.1838 | -9.9478
5.6525 | -34.1841 | -9.9583
5.6586 | -34.1844 | -9.9683
5.6648 | -34.1847 | -9.9787
5.6709 | -34.185 | -9.9892
5.677 | -34.1853 | -9.9996
5.6831 | -34.1856 | -10.01
5.6892 | -34.1859 | -10.0205
5.6953 | -34.1862 | -10.0308
5.7014 | -34.1865 | -10.0412
5.7075 | -34.1868 | -10.0515
5.7136 | -34.1871 | -10.0618
5.7197 | -34.1873 | -10.0722
5.7258 | -34.1876 | -10.0822
5.7319 | -34.1879 | -10.0926
5.738 | -34.1882 | -10.1029
5.7441 | -34.1886 | -10.1132
5.7502 | -34.1889 | -10.1233
5.7563 | -34.1892 | -10.1336
5.7624 | -34.1895 | -10.1442
5.7685 | -34.1898 | -10.1545
5.7746 | -34.1901 | -10.1651
5.7807 | -34.1903 | -10.1755
5.7868 | -34.1906 | -10.186
5.7929 | -34.1909 | -10.1963
5.799 | -34.1912 | -10.2064
5.8052 | -34.1915 | -10.2166
5.8113 | -34.1918 | -10.2268
5.8174 | -34.192 | -10.237
5.8235 | -34.1923 | -10.2472
5.8296 | -34.1926 | -10.2577
5.8357 | -34.1929 | -10.2679
5.8418 | -34.1932 | -10.2783
5.8479 | -34.1934 | -10.2887
5.854 | -34.1937 | -10.2992
5.8601 | -34.1941 | -10.3094
5.8662 | -34.1943 | -10.3199
5.8723 | -34.1946 | -10.33
5.8784 | -34.1949 | -10.3402
5.8845 | -34.1952 | -10.3504
5.8906 | -34.1955 | -10.3608
5.8967 | -34.1957 | -10.3709
5.9028 | -34.1961 | -10.381
5.9089 | -34.1964 | -10.3912
5.915 | -34.1967 | -10.4013
5.9211 | -34.197 | -10.4115
5.9272 | -34.1973 | -10.4215
5.9333 | -34.1976 | -10.4316
5.9394 | -34.1979 | -10.4418
5.9455 | -34.1982 | -10.4519
5.9517 | -34.1986 | -10.462
5.9578 | -34.1989 | -10.4724
5.9639 | -34.1992 | -10.4823
5.97 | -34.1995 | -10.4927
5.9761 | -34.1999 | -10.5028
5.9822 | -34.2001 | -10.5132
5.9883 | -34.2004 | -10.5235
5.9944 | -34.2008 | -10.5338
6.0005 | -34.2011 | -10.5443
6.0066 | -34.2014 | -10.5543
6.0127 | -34.2017 | -10.5642
6.0188 | -34.202 | -10.5747
6.0249 | -34.2023 | -10.585
6.031 | -34.2026 | -10.595
6.0371 | -34.2029 | -10.6053
6.0432 | -34.2032 | -10.6154
6.0493 | -34.2035 | -10.6255
6.0554 | -34.2038 | -10.6356
6.0615 | -34.2041 | -10.6458
6.0676 | -34.2044 | -10.6562
6.0737 | -34.2048 | -10.6666
6.0798 | -34.205 | -10.6768
6.0859 | -34.2054 | -10.687
6.0921 | -34.2057 | -10.6969
6.0982 | -34.206 | -10.7071
6.1043 | -34.2063 | -10.7172
6.1104 | -34.2066 | -10.7276
6.1165 | -34.207 | -10.7374
6.1226 | -34.2073 | -10.7475
6.1287 | -34.2076 | -10.7577
6.1348 | -34.2079 | -10.7679
6.1409 | -34.2082 | -10.778
6.147 | -34.2085 | -10.7881
6.1531 | -34.2088 | -10.7983
6.1592 | -34.2091 | -10.8083
6.1653 | -34.2094 | -10.8184
6.1714 | -34.2097 | -10.8286
6.1775 | -34.2101 | -10.8389
6.1836 | -34.2103 | -10.8491
6.1897 | -34.2107 | -10.8593
6.1958 | -34.211 | -10.8697
6.2019 | -34.2113 | -10.88
6.208 | -34.2116 | -10.8901
6.2141 | -34.212 | -10.9002
6.2202 | -34.2123 | -10.9105
6.2263 | -34.2126 | -10.9207
6.2325 | -34.2129 | -10.9306
6.2386 | -34.2132 | -10.9404
6.2447 | -34.2135 | -10.9508
6.2508 | -34.2139 | -10.961
6.2569 | -34.2142 | -10.9711
6.263 | -34.2145 | -10.9813
6.2691 | -34.2148 | -10.9913
6.2752 | -34.2151 | -11.0015
6.2813 | -34.2154 | -11.0117
6.2874 | -34.2157 | -11.0217
6.2935 | -34.216 | -11.0318
6.2996 | -34.2163 | -11.0419
6.3057 | -34.2166 | -11.0521
6.3118 | -34.2169 | -11.0623
6.3179 | -34.2172 | -11.0725
6.324 | -34.2175 | -11.0827
6.3301 | -34.2178 | -11.0928
6.3362 | -34.2181 | -11.1029
6.3423 | -34.2185 | -11.1132
6.3484 | -34.2188 | -11.1234
6.3545 | -34.2191 | -11.1333
6.3606 | -34.2194 | -11.1431
6.3667 | -34.2197 | -11.1532
6.3728 | -34.22 | -11.1633
6.379 | -34.2203 | -11.1735
6.3851 | -34.2206 | -11.1833
6.3912 | -34.2209 | -11.1934
6.3973 | -34.2212 | -11.2036
6.4034 | -34.2215 | -11.2135
6.4095 | -34.2218 | -11.2236
6.4156 | -34.2221 | -11.2338
6.4217 | -34.2224 | -11.2438
6.4278 | -34.2227 | -11.254
6.4339 | -34.223 | -11.264
6.44 | -34.2234 | -11.274
6.4461 | -34.2236 | -11.2843
6.4522 | -34.224 | -11.2942
6.4583 | -34.2242 | -11.3044
6.4644 | -34.2245 | -11.3146
6.4705 | -34.2248 | -11.3245
6.4766 | -34.2251 | -11.3346
6.4827 | -34.2254 | -11.3447
6.4888 | -34.2257 | -11.3547
6.4949 | -34.226 | -11.3645
6.501 | -34.2263 | -11.3747
6.5071 | -34.2267 | -11.3846
6.5132 | -34.227 | -11.3947
6.5194 | -34.2273 | -11.4048
6.5255 | -34.2276 | -11.4147
6.5316 | -34.2279 | -11.4249
6.5377 | -34.2282 | -11.435
6.5438 | -34.2285 | -11.4449
6.5499 | -34.2288 | -11.4549
6.556 | -34.2291 | -11.4649
6.5621 | -34.2294 | -11.4752
6.5682 | -34.2297 | -11.4853
6.5743 | -34.23 | -11.4954
6.5804 | -34.2303 | -11.5056
6.5865 | -34.2306 | -11.5157
6.5926 | -34.2309 | -11.5257
6.5987 | -34.2312 | -11.5354
6.6048 | -34.2316 | -11.5456
6.6109 | -34.2319 | -11.5557
6.617 | -34.2322 | -11.566
6.6231 | -34.2325 | -11.576
6.6292 | -34.2328 | -11.5861
6.6353 | -34.2331 | -11.5965
6.6414 | -34.2334 | -11.6064
6.6475 | -34.2337 | -11.6165
6.6536 | -34.234 | -11.6266
6.6597 | -34.2343 | -11.6364
6.6659 | -34.2346 | -11.6465
6.672 | -34.2349 | -11.6564
6.6781 | -34.2352 | -11.6666
6.6842 | -34.2356 | -11.6765
6.6903 | -34.2359 | -11.6864
6.6964 | -34.2362 | -11.6965
6.7025 | -34.2365 | -11.7063
6.7086 | -34.2368 | -11.7164
6.7147 | -34.2371 | -11.7265
6.7208 | -34.2374 | -11.7367
6.7269 | -34.2377 | -11.7466
6.733 | -34.238 | -11.7564
6.7391 | -34.2383 | -11.7665
6.7452 | -34.2386 | -11.7763
6.7513 | -34.2389 | -11.7861
6.7574 | -34.2393 | -11.7962
6.7635 | -34.2396 | -11.8061
6.7696 | -34.2399 | -11.816
6.7757 | -34.2402 | -11.826
6.7818 | -34.2405 | -11.8361
6.7879 | -34.2409 | -11.846
6.794 | -34.2412 | -11.8558
6.8001 | -34.2415 | -11.8658
6.8063 | -34.2418 | -11.8758
6.8124 | -34.2421 | -11.886
6.8185 | -34.2424 | -11.8962
6.8246 | -34.2426 | -11.9061
6.8307 | -34.243 | -11.916
6.8368 | -34.2433 | -11.9262
6.8429 | -34.2436 | -11.9363
6.849 | -34.2439 | -11.9463
6.8551 | -34.2443 | -11.9562
6.8612 | -34.2446 | -11.9662
6.8673 | -34.2448 | -11.9762
6.8734 | -34.2452 | -11.9864
6.8795 | -34.2455 | -11.9963
6.8856 | -34.2458 | -12.0063
6.8917 | -34.2461 | -12.0163
6.8978 | -34.2464 | -12.0261
6.9039 | -34.2467 | -12.0358
6.91 | -34.2471 | -12.0459
6.9161 | -34.2474 | -12.0556
6.9222 | -34.2477 | -12.0656
6.9283 | -34.248 | -12.0753
6.9344 | -34.2483 | -12.0853
6.9405 | -34.2487 | -12.0952
6.9466 | -34.249 | -12.1054
6.9528 | -34.2494 | -12.1155
6.9589 | -34.2497 | -12.1256
6.965 | -34.25 | -12.1357
6.9711 | -34.2504 | -12.1458
6.9772 | -34.2507 | -12.1557
6.9833 | -34.251 | -12.1657
6.9894 | -34.2513 | -12.1759
6.9955 | -34.2517 | -12.1856
7.0016 | -34.252 | -12.1956
7.0077 | -34.2523 | -12.2055
7.0138 | -34.2527 | -12.2154
7.0199 | -34.2529 | -12.2253
7.026 | -34.2533 | -12.2352
7.0321 | -34.2536 | -12.245
7.0382 | -34.2539 | -12.2547
7.0443 | -34.2543 | -12.2647
7.0504 | -34.2546 | -12.2747
7.0565 | -34.2549 | -12.2847
7.0626 | -34.2552 | -12.2944
7.0687 | -34.2556 | -12.3039
7.0748 | -34.2559 | -12.3136
7.0809 | -34.2562 | -12.3235
7.087 | -34.2566 | -12.3335
7.0932 | -34.2569 | -12.3435
7.0993 | -34.2572 | -12.3534
7.1054 | -34.2575 | -12.3632
7.1115 | -34.2579 | -12.3729
7.1176 | -34.2582 | -12.3827
7.1237 | -34.2585 | -12.3926
7.1298 | -34.2588 | -12.4024
7.1359 | -34.2591 | -12.4123
7.142 | -34.2595 | -12.4219
7.1481 | -34.2598 | -12.4317
7.1542 | -34.2601 | -12.4416
7.1603 | -34.2604 | -12.4515
7.1664 | -34.2608 | -12.4613
7.1725 | -34.2611 | -12.4712
7.1786 | -34.2615 | -12.481
7.1847 | -34.2618 | -12.491
7.1908 | -34.2621 | -12.501
7.1969 | -34.2624 | -12.511
7.203 | -34.2627 | -12.5207
7.2091 | -34.2631 | -12.5304
7.2152 | -34.2634 | -12.5403
7.2213 | -34.2637 | -12.5499
7.2274 | -34.2641 | -12.5597
7.2335 | -34.2644 | -12.5695
7.2397 | -34.2647 | -12.5793
7.2458 | -34.2651 | -12.5888
7.2519 | -34.2655 | -12.5987
7.258 | -34.2658 | -12.6087
7.2641 | -34.2661 | -12.6185
7.2702 | -34.2664 | -12.6286
7.2763 | -34.2668 | -12.6384
7.2824 | -34.2671 | -12.6484
7.2885 | -34.2674 | -12.6582
7.2946 | -34.2678 | -12.668
7.3007 | -34.2681 | -12.6779
7.3068 | -34.2684 | -12.688
7.3129 | -34.2687 | -12.6979
7.319 | -34.2691 | -12.7077
7.3251 | -34.2694 | -12.7173
7.3312 | -34.2697 | -12.7272
7.3373 | -34.2701 | -12.7368
7.3434 | -34.2704 | -12.7465
7.3495 | -34.2708 | -12.7562
7.3556 | -34.2711 | -12.7661
7.3617 | -34.2715 | -12.7761
7.3678 | -34.2718 | -12.7858
7.3739 | -34.2721 | -12.7958
7.3801 | -34.2725 | -12.8059
7.3862 | -34.2728 | -12.8156
7.3923 | -34.2732 | -12.8256
7.3984 | -34.2735 | -12.8352
7.4045 | -34.2739 | -12.845
7.4106 | -34.2742 | -12.8547
7.4167 | -34.2745 | -12.8646
7.4228 | -34.2749 | -12.8745
7.4289 | -34.2752 | -12.8843
7.435 | -34.2756 | -12.8939
7.4411 | -34.2759 | -12.9037
7.4472 | -34.2762 | -12.9136
7.4533 | -34.2766 | -12.9231
7.4594 | -34.277 | -12.9331
7.4655 | -34.2773 | -12.9428
7.4716 | -34.2776 | -12.9528
7.4777 | -34.2779 | -12.9625
7.4838 | -34.2783 | -12.9721
7.4899 | -34.2786 | -12.9818
7.496 | -34.279 | -12.9914
7.5021 | -34.2793 | -13.0011
7.5082 | -34.2796 | -13.0109
7.5143 | -34.28 | -13.0208
7.5204 | -34.2803 | -13.0307
7.5266 | -34.2806 | -13.0403
7.5327 | -34.281 | -13.05
7.5388 | -34.2813 | -13.06
7.5449 | -34.2817 | -13.07
7.551 | -34.282 | -13.0796
7.5571 | -34.2824 | -13.0895
7.5632 | -34.2827 | -13.0993
7.5693 | -34.2831 | -13.109
7.5754 | -34.2834 | -13.1186
7.5815 | -34.2837 | -13.1284
7.5876 | -34.2841 | -13.1384
7.5937 | -34.2845 | -13.1483
7.5998 | -34.2848 | -13.1581
7.6059 | -34.2852 | -13.1681
7.612 | -34.2855 | -13.1779
7.6181 | -34.2859 | -13.1875
7.6242 | -34.2862 | -13.1976
7.6303 | -34.2866 | -13.2075
7.6364 | -34.2869 | -13.2174
7.6425 | -34.2872 | -13.2273
7.6486 | -34.2875 | -13.2371
7.6547 | -34.2879 | -13.2468
7.6608 | -34.2883 | -13.2565
7.667 | -34.2886 | -13.2664
7.6731 | -34.2889 | -13.2762
7.6792 | -34.2893 | -13.2863
7.6853 | -34.2896 | -13.2962
7.6914 | -34.29 | -13.3061
7.6975 | -34.2903 | -13.316
7.7036 | -34.2907 | -13.3258
7.7097 | -34.291 | -13.3357
7.7158 | -34.2914 | -13.3456
7.7219 | -34.2917 | -13.3552
7.728 | -34.2921 | -13.3651
7.7341 | -34.2924 | -13.375
7.7402 | -34.2928 | -13.3848
7.7463 | -34.2931 | -13.3947
7.7524 | -34.2935 | -13.4048
7.7585 | -34.2938 | -13.4146
7.7646 | -34.2942 | -13.4244
7.7707 | -34.2945 | -13.4343
7.7768 | -34.2948 | -13.4443
7.7829 | -34.2952 | -13.4544
7.789 | -34.2955 | -13.4644
7.7951 | -34.2958 | -13.4742
7.8012 | -34.2962 | -13.4841
7.8073 | -34.2966 | -13.4941
7.8135 | -34.2969 | -13.5039
7.8196 | -34.2973 | -13.5139
7.8257 | -34.2976 | -13.524
7.8318 | -34.2979 | -13.5339
7.8379 | -34.2983 | -13.5436
7.844 | -34.2987 | -13.5534
7.8501 | -34.299 | -13.5635
7.8562 | -34.2993 | -13.5736
7.8623 | -34.2997 | -13.5833
7.8684 | -34.3 | -13.5932
7.8745 | -34.3004 | -13.6032
7.8806 | -34.3007 | -13.6132
7.8867 | -34.3011 | -13.6228
7.8928 | -34.3014 | -13.6325
7.8989 | -34.3018 | -13.6424
7.905 | -34.3021 | -13.6523
7.9111 | -34.3024 | -13.6622
7.9172 | -34.3028 | -13.6721
7.9233 | -34.3031 | -13.682
7.9294 | -34.3035 | -13.692
7.9355 | -34.3038 | -13.7017
7.9416 | -34.3041 | -13.7116
7.9477 | -34.3045 | -13.7213
7.9539 | -34.3048 | -13.7311
7.96 | -34.3052 | -13.741
7.9661 | -34.3055 | -13.7509
7.9722 | -34.3059 | -13.7606
7.9783 | -34.3062 | -13.7707
7.9844 | -34.3066 | -13.7804
7.9905 | -34.3069 | -13.7905
7.9966 | -34.3073 | -13.8005
8.0027 | -34.3076 | -13.8103
8.0088 | -34.3079 | -13.8202
8.0149 | -34.3083 | -13.8301
8.021 | -34.3086 | -13.84
8.0271 | -34.309 | -13.8501
8.0332 | -34.3093 | -13.8599
8.0393 | -34.3097 | -13.8699
8.0454 | -34.31 | -13.88
8.0515 | -34.3104 | -13.8899
8.0576 | -34.3107 | -13.9
8.0637 | -34.3111 | -13.9097
8.0698 | -34.3114 | -13.9195
8.0759 | -34.3117 | -13.9293
8.082 | -34.3121 | -13.9391
8.0881 | -34.3124 | -13.949
8.0942 | -34.3128 | -13.9589
8.1004 | -34.3132 | -13.9689
8.1065 | -34.3135 | -13.9786
8.1126 | -34.3138 | -13.9886
8.1187 | -34.3142 | -13.9986
8.1248 | -34.3145 | -14.0083
8.1309 | -34.3149 | -14.018
8.137 | -34.3152 | -14.0277
8.1431 | -34.3156 | -14.0378
8.1492 | -34.3159 | -14.0476
8.1553 | -34.3162 | -14.0576
8.1614 | -34.3166 | -14.0676
8.1675 | -34.3169 | -14.0775
8.1736 | -34.3172 | -14.0875
8.1797 | -34.3176 | -14.0975
8.1858 | -34.3179 | -14.1074
8.1919 | -34.3183 | -14.1173
8.198 | -34.3186 | -14.127
8.2041 | -34.319 | -14.1369
8.2102 | -34.3193 | -14.1468
8.2163 | -34.3197 | -14.1566
8.2224 | -34.32 | -14.1664
8.2285 | -34.3204 | -14.1761
8.2346 | -34.3207 | -14.1859
8.2408 | -34.3211 | -14.1956
8.2469 | -34.3214 | -14.2056
8.253 | -34.3218 | -14.2153
8.2591 | -34.3221 | -14.2253
8.2652 | -34.3225 | -14.2352
8.2713 | -34.3228 | -14.2452
8.2774 | -34.3232 | -14.2551
8.2835 | -34.3235 | -14.2652
8.2896 | -34.3239 | -14.2749
8.2957 | -34.3242 | -14.2849
8.3018 | -34.3246 | -14.2948
8.3079 | -34.3249 | -14.3046
8.314 | -34.3253 | -14.3144
8.3201 | -34.3257 | -14.3242
8.3262 | -34.326 | -14.3342
8.3323 | -34.3264 | -14.344
8.3384 | -34.3267 | -14.3539
8.3445 | -34.327 | -14.3637
8.3506 | -34.3274 | -14.3736
8.3567 | -34.3278 | -14.3835
8.3628 | -34.3281 | -14.3934
8.3689 | -34.3284 | -14.4032
8.375 | -34.3288 | -14.4131
8.3812 | -34.3291 | -14.4228
8.3873 | -34.3295 | -14.4326
8.3934 | -34.3298 | -14.4426
8.3995 | -34.3302 | -14.4523
8.4056 | -34.3305 | -14.4621
8.4117 | -34.3309 | -14.4721
8.4178 | -34.3312 | -14.4819
8.4239 | -34.3315 | -14.4916
8.43 | -34.3319 | -14.5015
8.4361 | -34.3322 | -14.5115
8.4422 | -34.3326 | -14.5212
8.4483 | -34.3329 | -14.531
8.4544 | -34.3333 | -14.5408
8.4605 | -34.3336 | -14.5505
8.4666 | -34.3339 | -14.5605
8.4727 | -34.3343 | -14.5704
8.4788 | -34.3346 | -14.5802
8.4849 | -34.335 | -14.5901
8.491 | -34.3353 | -14.5999
8.4971 | -34.3357 | -14.6096
8.5032 | -34.336 | -14.6193
8.5093 | -34.3364 | -14.629
8.5154 | -34.3368 | -14.6389
8.5215 | -34.3371 | -14.6485
8.5277 | -34.3375 | -14.6584
8.5338 | -34.3378 | -14.6682
8.5399 | -34.3381 | -14.678
8.546 | -34.3385 | -14.6877
8.5521 | -34.3389 | -14.6975
8.5582 | -34.3392 | -14.7073
8.5643 | -34.3396 | -14.7168
8.5704 | -34.3399 | -14.7264
8.5765 | -34.3403 | -14.7363
8.5826 | -34.3406 | -14.7458
8.5887 | -34.341 | -14.7555
8.5948 | -34.3414 | -14.7654
8.6009 | -34.3417 | -14.775
8.607 | -34.3421 | -14.7848
8.6131 | -34.3424 | -14.7947
8.6192 | -34.3428 | -14.8043
8.6253 | -34.3432 | -14.814
8.6314 | -34.3435 | -14.8237
8.6375 | -34.3438 | -14.8335
8.6436 | -34.3442 | -14.8433
8.6497 | -34.3446 | -14.853
8.6558 | -34.3449 | -14.8627
8.6619 | -34.3452 | -14.8726
8.6681 | -34.3456 | -14.8825
8.6742 | -34.346 | -14.8922
8.6803 | -34.3463 | -14.9019
8.6864 | -34.3467 | -14.9114
8.6925 | -34.3471 | -14.9209
8.6986 | -34.3474 | -14.9307
8.7047 | -34.3478 | -14.9406
8.7108 | -34.3481 | -14.9504
8.7169 | -34.3485 | -14.9602
8.723 | -34.3488 | -14.9698
8.7291 | -34.3492 | -14.9794
8.7352 | -34.3495 | -14.9893
8.7413 | -34.3499 | -14.9991
8.7474 | -34.3502 | -15.0086
8.7535 | -34.3506 | -15.0184
8.7596 | -34.351 | -15.0282
8.7657 | -34.3513 | -15.038
8.7718 | -34.3517 | -15.0479
8.7779 | -34.352 | -15.0577
8.784 | -34.3524 | -15.0674
8.7901 | -34.3528 | -15.0772
8.7962 | -34.3531 | -15.0867
8.8023 | -34.3535 | -15.0964
8.8084 | -34.3538 | -15.106
8.8146 | -34.3542 | -15.1157
8.8207 | -34.3545 | -15.1254
8.8268 | -34.3549 | -15.1349
8.8329 | -34.3553 | -15.1446
8.839 | -34.3557 | -15.1543
8.8451 | -34.356 | -15.1639
8.8512 | -34.3564 | -15.1735
8.8573 | -34.3567 | -15.1832
8.8634 | -34.3571 | -15.1927
8.8695 | -34.3575 | -15.2026
8.8756 | -34.3578 | -15.2124
8.8817 | -34.3582 | -15.222
8.8878 | -34.3585 | -15.2318
8.8939 | -34.3589 | -15.2415
8.9 | -34.3593 | -15.2512
8.9061 | -34.3596 | -15.2608
8.9122 | -34.36 | -15.2705
8.9183 | -34.3604 | -15.2802
8.9244 | -34.3608 | -15.2901
8.9305 | -34.3612 | -15.2995
8.9366 | -34.3616 | -15.3092
8.9427 | -34.3619 | -15.3189
8.9488 | -34.3623 | -15.3286
8.955 | -34.3626 | -15.3382
8.9611 | -34.363 | -15.348
8.9672 | -34.3634 | -15.3577
8.9733 | -34.3638 | -15.3672
8.9794 | -34.3642 | -15.3768
8.9855 | -34.3646 | -15.3865
8.9916 | -34.3649 | -15.3964
8.9977 | -34.3653 | -15.406
9.0038 | -34.3657 | -15.4155
9.0099 | -34.3661 | -15.4252
9.016 | -34.3664 | -15.4349
9.0221 | -34.3668 | -15.4444
9.0282 | -34.3672 | -15.4538
9.0343 | -34.3676 | -15.4635
9.0404 | -34.368 | -15.4731
9.0465 | -34.3683 | -15.483
9.0526 | -34.3687 | -15.4925
9.0587 | -34.3691 | -15.5022
9.0648 | -34.3695 | -15.5117
9.0709 | -34.3699 | -15.5213
9.077 | -34.3703 | -15.531
9.0831 | -34.3707 | -15.5408
9.0892 | -34.371 | -15.5504
9.0953 | -34.3714 | -15.5602
9.1015 | -34.3718 | -15.5698
9.1076 | -34.3722 | -15.5793
9.1137 | -34.3726 | -15.5889
9.1198 | -34.373 | -15.5985
9.1259 | -34.3734 | -15.6081
9.132 | -34.3738 | -15.6179
9.1381 | -34.3742 | -15.6275
9.1442 | -34.3746 | -15.6368
9.1503 | -34.3749 | -15.6465
9.1564 | -34.3753 | -15.6561
9.1625 | -34.3757 | -15.6658
9.1686 | -34.3761 | -15.6753
9.1747 | -34.3765 | -15.685
9.1808 | -34.3769 | -15.6946
9.1869 | -34.3773 | -15.7042
9.193 | -34.3777 | -15.7138
9.1991 | -34.378 | -15.7235
9.2052 | -34.3785 | -15.7331
9.2113 | -34.3788 | -15.7427
9.2174 | -34.3793 | -15.7524
9.2235 | -34.3797 | -15.762
9.2296 | -34.3801 | -15.7716
9.2357 | -34.3805 | -15.7813
9.2419 | -34.3809 | -15.7909
9.248 | -34.3813 | -15.8005
9.2541 | -34.3817 | -15.8102
9.2602 | -34.3821 | -15.8199
9.2663 | -34.3825 | -15.8293
9.2724 | -34.3829 | -15.8389
9.2785 | -34.3833 | -15.8485
9.2846 | -34.3837 | -15.8579
9.2907 | -34.3841 | -15.8675
9.2968 | -34.3845 | -15.8773
9.3029 | -34.3849 | -15.887
9.309 | -34.3853 | -15.8966
9.3151 | -34.3857 | -15.906
9.3212 | -34.3861 | -15.9155
9.3273 | -34.3865 | -15.9251
9.3334 | -34.3869 | -15.9345
9.3395 | -34.3873 | -15.944
9.3456 | -34.3877 | -15.9534
9.3517 | -34.3881 | -15.9631
9.3578 | -34.3885 | -15.9726
9.3639 | -34.3889 | -15.9821
9.37 | -34.3893 | -15.9917
9.3761 | -34.3897 | -16.0014
9.3822 | -34.3901 | -16.011
9.3884 | -34.3905 | -16.0204
9.3945 | -34.3909 | -16.03
9.4006 | -34.3914 | -16.0396
9.4067 | -34.3918 | -16.0492
9.4128 | -34.3922 | -16.0588
9.4189 | -34.3926 | -16.0686
9.425 | -34.393 | -16.0781
9.4311 | -34.3934 | -16.0876
9.4372 | -34.3938 | -16.0971
9.4433 | -34.3942 | -16.1066
9.4494 | -34.3947 | -16.1161
9.4555 | -34.3951 | -16.1257
9.4616 | -34.3955 | -16.1352
9.4677 | -34.3959 | -16.1448
9.4738 | -34.3963 | -16.1544
9.4799 | -34.3967 | -16.164
9.486 | -34.3971 | -16.1735
9.4921 | -34.3975 | -16.183
9.4982 | -34.398 | -16.1924
9.5043 | -34.3984 | -16.2019
9.5104 | -34.3988 | -16.2113
9.5165 | -34.3992 | -16.2209
9.5226 | -34.3996 | -16.2306
9.5288 | -34.4 | -16.24
9.5349 | -34.4004 | -16.2496
9.541 | -34.4008 | -16.2589
9.5471 | -34.4012 | -16.2684
9.5532 | -34.4016 | -16.2781
9.5593 | -34.4021 | -16.2877
9.5654 | -34.4025 | -16.2972
9.5715 | -34.4029 | -16.3067
9.5776 | -34.4033 | -16.3164
9.5837 | -34.4037 | -16.3259
9.5898 | -34.404 | -16.3355
9.5959 | -34.4045 | -16.3451
9.602 | -34.4049 | -16.3544
9.6081 | -34.4053 | -16.3639
9.6142 | -34.4057 | -16.3734
9.6203 | -34.4061 | -16.383
9.6264 | -34.4065 | -16.3924
9.6325 | -34.4069 | -16.402
9.6386 | -34.4073 | -16.4114
9.6447 | -34.4077 | -16.421
9.6508 | -34.4081 | -16.4304
9.6569 | -34.4085 | -16.44
9.663 | -34.409 | -16.4496
9.6691 | -34.4094 | -16.4591
9.6753 | -34.4098 | -16.4685
9.6814 | -34.4102 | -16.4779
9.6875 | -34.4106 | -16.4872
9.6936 | -34.411 | -16.4965
9.6997 | -34.4115 | -16.5061
9.7058 | -34.4119 | -16.5156
9.7119 | -34.4124 | -16.5252
9.718 | -34.4128 | -16.5348
9.7241 | -34.4132 | -16.5442
9.7302 | -34.4136 | -16.5537
9.7363 | -34.414 | -16.5631
9.7424 | -34.4144 | -16.5725
9.7485 | -34.4148 | -16.582
9.7546 | -34.4152 | -16.5913
9.7607 | -34.4156 | -16.601
9.7668 | -34.4161 | -16.6105
9.7729 | -34.4165 | -16.62
9.779 | -34.4169 | -16.6292
9.7851 | -34.4173 | -16.6389
9.7912 | -34.4178 | -16.6483
9.7973 | -34.4182 | -16.6579
9.8034 | -34.4186 | -16.6671
9.8095 | -34.419 | -16.6768
9.8157 | -34.4195 | -16.6863
9.8218 | -34.4199 | -16.6959
9.8279 | -34.4203 | -16.7052
9.834 | -34.4207 | -16.7148
9.8401 | -34.4211 | -16.7243
9.8462 | -34.4216 | -16.7335
9.8523 | -34.422 | -16.743
9.8584 | -34.4224 | -16.7524
9.8645 | -34.4228 | -16.7621
9.8706 | -34.4232 | -16.7716
9.8767 | -34.4236 | -16.781
9.8828 | -34.424 | -16.7903
9.8889 | -34.4244 | -16.7996
9.895 | -34.4248 | -16.8091
9.9011 | -34.4252 | -16.8186
9.9072 | -34.4256 | -16.828
9.9133 | -34.426 | -16.8375
9.9194 | -34.4265 | -16.8468
9.9255 | -34.4269 | -16.8562
9.9316 | -34.4273 | -16.8656
9.9377 | -34.4277 | -16.8751
9.9438 | -34.4281 | -16.8846
9.9499 | -34.4285 | -16.8942
9.956 | -34.4289 | -16.9036
9.9622 | -34.4293 | -16.9132
9.9683 | -34.4298 | -16.9226
9.9744 | -34.4302 | -16.9322
9.9805 | -34.4306 | -16.9416
9.9866 | -34.431 | -16.951
9.9927 | -34.4314 | -16.9603
9.9988 | -34.4318 | -16.9699
10.0049 | -34.4322 | -16.9793
10.011 | -34.4327 | -16.9888
10.0171 | -34.4331 | -16.9981
10.0232 | -34.4335 | -17.0074
10.0293 | -34.4339 | -17.0169
10.0354 | -34.4344 | -17.0264
10.0415 | -34.4348 | -17.0359
10.0476 | -34.4352 | -17.0454
10.0537 | -34.4356 | -17.055
10.0598 | -34.436 | -17.0643
10.0659 | -34.4364 | -17.0738
10.072 | -34.4368 | -17.0831
10.0781 | -34.4372 | -17.0924
10.0842 | -34.4377 | -17.1019
10.0903 | -34.4381 | -17.1113
10.0964 | -34.4385 | -17.1204
10.1026 | -34.4389 | -17.1299
10.1087 | -34.4393 | -17.1393
10.1148 | -34.4397 | -17.1488
10.1209 | -34.4401 | -17.1581
10.127 | -34.4405 | -17.1674
10.1331 | -34.4409 | -17.1767
10.1392 | -34.4413 | -17.1859
10.1453 | -34.4418 | -17.1952
10.1514 | -34.4422 | -17.2046
10.1575 | -34.4426 | -17.2142
10.1636 | -34.443 | -17.2233
10.1697 | -34.4434 | -17.2328
10.1758 | -34.4438 | -17.242
10.1819 | -34.4443 | -17.2513
10.188 | -34.4446 | -17.2607
10.1941 | -34.445 | -17.2702
10.2002 | -34.4454 | -17.2795
10.2063 | -34.4458 | -17.2889
10.2124 | -34.4463 | -17.2983
10.2185 | -34.4467 | -17.3078
10.2246 | -34.4471 | -17.317
10.2307 | -34.4475 | -17.3265
10.2368 | -34.4479 | -17.3355
10.2429 | -34.4483 | -17.3449
10.2491 | -34.4487 | -17.3544
10.2552 | -34.4491 | -17.3638
10.2613 | -34.4495 | -17.3733
10.2674 | -34.45 | -17.3827
10.2735 | -34.4504 | -17.3923
10.2796 | -34.4508 | -17.4016
10.2857 | -34.4512 | -17.4111
10.2918 | -34.4516 | -17.4206
10.2979 | -34.452 | -17.4299
10.304 | -34.4525 | -17.4391
10.3101 | -34.4529 | -17.4485
10.3162 | -34.4533 | -17.458
10.3223 | -34.4537 | -17.4672
10.3284 | -34.4541 | -17.4766
10.3345 | -34.4545 | -17.4858
10.3406 | -34.4549 | -17.4951
10.3467 | -34.4553 | -17.5044
10.3528 | -34.4557 | -17.5136
10.3589 | -34.4561 | -17.523
10.365 | -34.4565 | -17.5325
10.3711 | -34.4569 | -17.5419
10.3772 | -34.4574 | -17.5513
10.3833 | -34.4578 | -17.5607
10.3895 | -34.4582 | -17.57
10.3956 | -34.4586 | -17.5794
10.4017 | -34.459 | -17.5885
10.4078 | -34.4594 | -17.5979
10.4139 | -34.4598 | -17.6072
10.42 | -34.4603 | -17.6167
10.4261 | -34.4607 | -17.626
10.4322 | -34.4611 | -17.6353
10.4383 | -34.4615 | -17.6446
10.4444 | -34.4619 | -17.6538
10.4505 | -34.4623 | -17.6632
10.4566 | -34.4627 | -17.6724
10.4627 | -34.4631 | -17.6816
10.4688 | -34.4635 | -17.6908
10.4749 | -34.4639 | -17.7003
10.481 | -34.4643 | -17.7096
10.4871 | -34.4647 | -17.7188
10.4932 | -34.4652 | -17.728
10.4993 | -34.4656 | -17.7374
10.5054 | -34.466 | -17.7467
10.5115 | -34.4664 | -17.756
10.5176 | -34.4668 | -17.7653
10.5237 | -34.4672 | -17.7747
10.5298 | -34.4676 | -17.7839
10.536 | -34.468 | -17.7933
10.5421 | -34.4684 | -17.8024
10.5482 | -34.4688 | -17.8118
10.5543 | -34.4693 | -17.821
10.5604 | -34.4697 | -17.8301
10.5665 | -34.4701 | -17.8392
10.5726 | -34.4705 | -17.8485
10.5787 | -34.4709 | -17.8578
10.5848 | -34.4713 | -17.8667
10.5909 | -34.4718 | -17.8761
10.597 | -34.4721 | -17.8855
10.6031 | -34.4726 | -17.895
10.6092 | -34.473 | -17.9042
10.6153 | -34.4734 | -17.9136
10.6214 | -34.4738 | -17.9227
10.6275 | -34.4742 | -17.932
10.6336 | -34.4746 | -17.9413
10.6397 | -34.475 | -17.9506
10.6458 | -34.4754 | -17.9598
10.6519 | -34.4758 | -17.9692
10.658 | -34.4762 | -17.9783
10.6641 | -34.4766 | -17.9878
10.6702 | -34.477 | -17.9971
10.6764 | -34.4775 | -18.0064
10.6825 | -34.4779 | -18.0155
10.6886 | -34.4783 | -18.0246
10.6947 | -34.4787 | -18.0339
10.7008 | -34.4791 | -18.0431
10.7069 | -34.4796 | -18.0525
10.713 | -34.48 | -18.0617
10.7191 | -34.4804 | -18.0709
10.7252 | -34.4808 | -18.0799
10.7313 | -34.4812 | -18.089
10.7374 | -34.4816 | -18.0983
10.7435 | -34.482 | -18.1076
10.7496 | -34.4824 | -18.1167
10.7557 | -34.4828 | -18.126
10.7618 | -34.4832 | -18.135
10.7679 | -34.4836 | -18.1443
10.774 | -34.484 | -18.1536
10.7801 | -34.4844 | -18.1628
10.7862 | -34.4848 | -18.1719
10.7923 | -34.4852 | -18.1812
10.7984 | -34.4857 | -18.1903
10.8045 | -34.4861 | -18.1993
10.8106 | -34.4865 | -18.2086
10.8168 | -34.4869 | -18.2177
10.8229 | -34.4873 | -18.2269
10.829 | -34.4877 | -18.2361
10.8351 | -34.4881 | -18.245
10.8412 | -34.4886 | -18.2542
10.8473 | -34.489 | -18.2632
10.8534 | -34.4894 | -18.2724
10.8595 | -34.4898 | -18.2816
10.8656 | -34.4903 | -18.2906
10.8717 | -34.4907 | -18.2998
10.8778 | -34.4911 | -18.309
10.8839 | -34.4915 | -18.3181
10.89 | -34.4919 | -18.3272
10.8961 | -34.4924 | -18.3363
10.9022 | -34.4928 | -18.3453
10.9083 | -34.4932 | -18.3544
10.9144 | -34.4936 | -18.3635
10.9205 | -34.494 | -18.3726
10.9266 | -34.4944 | -18.3817
10.9327 | -34.4948 | -18.3907
10.9388 | -34.4953 | -18.3997
10.9449 | -34.4957 | -18.4089
10.951 | -34.4961 | -18.418
10.9571 | -34.4965 | -18.4272
10.9633 | -34.4969 | -18.4364
10.9694 | -34.4973 | -18.4457
10.9755 | -34.4978 | -18.4548
10.9816 | -34.4982 | -18.464
10.9877 | -34.4986 | -18.473
10.9938 | -34.499 | -18.4824
10.9999 | -34.4994 | -18.4913
11.006 | -34.4998 | -18.5002
11.0121 | -34.5003 | -18.5094
11.0182 | -34.5007 | -18.5185
11.0243 | -34.5011 | -18.5276
11.0304 | -34.5015 | -18.5368
11.0365 | -34.5019 | -18.5458
11.0426 | -34.5024 | -18.5548
11.0487 | -34.5028 | -18.5639
11.0548 | -34.5032 | -18.573
11.0609 | -34.5036 | -18.5821
11.067 | -34.5041 | -18.5912
11.0731 | -34.5045 | -18.6002
11.0792 | -34.5049 | -18.6093
11.0853 | -34.5054 | -18.6185
11.0914 | -34.5058 | -18.6277
11.0975 | -34.5062 | -18.6368
11.1037 | -34.5066 | -18.6459
11.1098 | -34.507 | -18.6552
11.1159 | -34.5075 | -18.664
11.122 | -34.5079 | -18.6729
11.1281 | -34.5083 | -18.682
11.1342 | -34.5087 | -18.6911
11.1403 | -34.5091 | -18.7003
11.1464 | -34.5095 | -18.7095
11.1525 | -34.51 | -18.7183
11.1586 | -34.5104 | -18.7275
11.1647 | -34.5108 | -18.7365
11.1708 | -34.5112 | -18.7455
11.1769 | -34.5116 | -18.7546
11.183 | -34.512 | -18.7635
11.1891 | -34.5125 | -18.7726
11.1952 | -34.5129 | -18.7818
11.2013 | -34.5133 | -18.7909
11.2074 | -34.5138 | -18.7999
11.2135 | -34.5142 | -18.8091
11.2196 | -34.5146 | -18.818
11.2257 | -34.515 | -18.8272
11.2318 | -34.5155 | -18.8363
11.2379 | -34.5159 | -18.8454
11.244 | -34.5164 | -18.8545
11.2502 | -34.5168 | -18.8635
11.2563 | -34.5172 | -18.8725
11.2624 | -34.5177 | -18.8814
11.2685 | -34.5181 | -18.8904
11.2746 | -34.5185 | -18.8994
11.2807 | -34.5189 | -18.9085
11.2868 | -34.5194 | -18.9176
11.2929 | -34.5198 | -18.9263
11.299 | -34.5202 | -18.9354
11.3051 | -34.5207 | -18.9445
11.3112 | -34.5211 | -18.9536
11.3173 | -34.5215 | -18.9623
11.3234 | -34.5219 | -18.9711
11.3295 | -34.5224 | -18.9801
11.3356 | -34.5228 | -18.9893
11.3417 | -34.5232 | -18.9981
11.3478 | -34.5236 | -19.0074
11.3539 | -34.524 | -19.0166
11.36 | -34.5245 | -19.0256
11.3661 | -34.5249 | -19.0347
11.3722 | -34.5253 | -19.0435
11.3783 | -34.5258 | -19.0525
11.3844 | -34.5262 | -19.0616
11.3906 | -34.5267 | -19.0705
11.3967 | -34.5271 | -19.0792
11.4028 | -34.5275 | -19.0885
11.4089 | -34.528 | -19.0974
11.415 | -34.5284 | -19.1064
11.4211 | -34.5288 | -19.1155
11.4272 | -34.5292 | -19.1245
11.4333 | -34.5296 | -19.1332
11.4394 | -34.53 | -19.1423
11.4455 | -34.5304 | -19.1512
11.4516 | -34.5309 | -19.1603
11.4577 | -34.5313 | -19.1692
11.4638 | -34.5318 | -19.1779
11.4699 | -34.5322 | -19.187
11.476 | -34.5326 | -19.1961
11.4821 | -34.5331 | -19.205
11.4882 | -34.5335 | -19.2142
11.4943 | -34.5339 | -19.2232
11.5004 | -34.5344 | -19.232
11.5065 | -34.5348 | -19.241
11.5126 | -34.5352 | -19.2498
11.5187 | -34.5357 | -19.2588
11.5248 | -34.5361 | -19.2677
11.5309 | -34.5366 | -19.2765
11.5371 | -34.537 | -19.2853
11.5432 | -34.5374 | -19.2942
11.5493 | -34.5378 | -19.3033
11.5554 | -34.5382 | -19.3121
11.5615 | -34.5386 | -19.3213
11.5676 | -34.5391 | -19.3302
11.5737 | -34.5395 | -19.3389
11.5798 | -34.5399 | -19.3475
11.5859 | -34.5404 | -19.3565
11.592 | -34.5407 | -19.3654
11.5981 | -34.5412 | -19.3743
11.6042 | -34.5416 | -19.3831
11.6103 | -34.542 | -19.3919
11.6164 | -34.5425 | -19.4006
11.6225 | -34.5429 | -19.4094
11.6286 | -34.5433 | -19.4182
11.6347 | -34.5437 | -19.4271
11.6408 | -34.5442 | -19.4358
11.6469 | -34.5446 | -19.4448
11.653 | -34.545 | -19.4536
11.6591 | -34.5454 | -19.4625
11.6652 | -34.5459 | -19.4712
11.6713 | -34.5463 | -19.4801
11.6775 | -34.5468 | -19.4887
11.6836 | -34.5472 | -19.4975
11.6897 | -34.5477 | -19.5064
11.6958 | -34.5481 | -19.5154
11.7019 | -34.5485 | -19.5244
11.708 | -34.549 | -19.5331
11.7141 | -34.5494 | -19.5418
11.7202 | -34.5498 | -19.5506
11.7263 | -34.5503 | -19.5595
11.7324 | -34.5507 | -19.5685
11.7385 | -34.5511 | -19.5773
11.7446 | -34.5515 | -19.586
11.7507 | -34.5519 | -19.5944
11.7568 | -34.5524 | -19.6032
11.7629 | -34.5528 | -19.6121
11.769 | -34.5532 | -19.6209
11.7751 | -34.5536 | -19.6297
11.7812 | -34.5541 | -19.6382
11.7873 | -34.5545 | -19.6469
11.7934 | -34.5549 | -19.6558
11.7995 | -34.5553 | -19.6646
11.8056 | -34.5558 | -19.6733
11.8117 | -34.5562 | -19.6822
11.8178 | -34.5567 | -19.6908
11.824 | -34.5571 | -19.6995
11.8301 | -34.5575 | -19.7083
11.8362 | -34.5579 | -19.717
11.8423 | -34.5584 | -19.7258
11.8484 | -34.5588 | -19.7347
11.8545 | -34.5593 | -19.7431
11.8606 | -34.5597 | -19.7517
11.8667 | -34.5601 | -19.7605
11.8728 | -34.5606 | -19.7692
11.8789 | -34.561 | -19.7779
11.885 | -34.5615 | -19.7866
11.8911 | -34.5619 | -19.7953
11.8972 | -34.5624 | -19.804
11.9033 | -34.5628 | -19.8126
11.9094 | -34.5632 | -19.8213
11.9155 | -34.5636 | -19.83
11.9216 | -34.564 | -19.8385
11.9277 | -34.5645 | -19.8472
11.9338 | -34.5649 | -19.8561
11.9399 | -34.5653 | -19.8651
11.946 | -34.5658 | -19.8738
11.9521 | -34.5662 | -19.8824
11.9582 | -34.5667 | -19.8911
11.9644 | -34.5671 | -19.8998
11.9705 | -34.5675 | -19.9085
11.9766 | -34.568 | -19.9171
11.9827 | -34.5684 | -19.9258
11.9888 | -34.5689 | -19.9345
11.9949 | -34.5694 | -19.9427
12.001 | -34.5698 | -19.9515
12.0071 | -34.5702 | -19.9601
12.0132 | -34.5707 | -19.9688
12.0193 | -34.5711 | -19.9776
12.0254 | -34.5715 | -19.9863
12.0315 | -34.572 | -19.9947
12.0376 | -34.5724 | -20.0034
12.0437 | -34.5728 | -20.0122
12.0498 | -34.5733 | -20.0209
12.0559 | -34.5737 | -20.0295
12.062 | -34.5742 | -20.0378
12.0681 | -34.5746 | -20.0465
12.0742 | -34.5751 | -20.0551
12.0803 | -34.5755 | -20.0638
12.0864 | -34.5759 | -20.0725
12.0925 | -34.5764 | -20.0812
12.0986 | -34.5768 | -20.0897
12.1047 | -34.5773 | -20.0983
12.1109 | -34.5777 | -20.1067
12.117 | -34.5781 | -20.1153
12.1231 | -34.5786 | -20.1239
12.1292 | -34.579 | -20.1325
12.1353 | -34.5795 | -20.141
12.1414 | -34.5799 | -20.1497
12.1475 | -34.5804 | -20.1584
12.1536 | -34.5808 | -20.1671
12.1597 | -34.5812 | -20.1757
12.1658 | -34.5817 | -20.1842
12.1719 | -34.5821 | -20.1925
12.178 | -34.5825 | -20.2011
12.1841 | -34.583 | -20.2098
12.1902 | -34.5834 | -20.2184
12.1963 | -34.5838 | -20.2271
12.2024 | -34.5843 | -20.2354
12.2085 | -34.5847 | -20.2439
12.2146 | -34.5852 | -20.2526
12.2207 | -34.5856 | -20.2612
12.2268 | -34.5861 | -20.2699
12.2329 | -34.5865 | -20.2784
12.239 | -34.587 | -20.2866
12.2451 | -34.5874 | -20.2951
12.2513 | -34.5878 | -20.3038
12.2574 | -34.5883 | -20.3124
12.2635 | -34.5887 | -20.321
12.2696 | -34.5892 | -20.3294
12.2757 | -34.5896 | -20.3376
12.2818 | -34.5901 | -20.3462
12.2879 | -34.5905 | -20.3549
12.294 | -34.5909 | -20.3632
12.3001 | -34.5914 | -20.3717
12.3062 | -34.5919 | -20.3803
12.3123 | -34.5923 | -20.3886
12.3184 | -34.5928 | -20.397
12.3245 | -34.5932 | -20.4054
12.3306 | -34.5936 | -20.4139
12.3367 | -34.5941 | -20.4225
12.3428 | -34.5945 | -20.431
12.3489 | -34.595 | -20.4391
12.355 | -34.5954 | -20.4478
12.3611 | -34.5958 | -20.4565
12.3672 | -34.5963 | -20.4649
12.3733 | -34.5967 | -20.4734
12.3794 | -34.5972 | -20.4816
12.3855 | -34.5976 | -20.49
12.3916 | -34.598 | -20.4985
12.3978 | -34.5985 | -20.507
12.4039 | -34.5989 | -20.5155
12.41 | -34.5994 | -20.5241
12.4161 | -34.5999 | -20.5322
12.4222 | -34.6003 | -20.5407
12.4283 | -34.6008 | -20.5491
12.4344 | -34.6012 | -20.5577
12.4405 | -34.6016 | -20.5663
12.4466 | -34.6021 | -20.5747
12.4527 | -34.6025 | -20.583
12.4588 | -34.603 | -20.5915
12.4649 | -34.6034 | -20.6
12.471 | -34.6039 | -20.6086
12.4771 | -34.6043 | -20.617
12.4832 | -34.6047 | -20.6254
12.4893 | -34.6052 | -20.6337
12.4954 | -34.6056 | -20.642
12.5015 | -34.6061 | -20.6504
12.5076 | -34.6065 | -20.6589
12.5137 | -34.607 | -20.6674
12.5198 | -34.6074 | -20.6756
12.5259 | -34.6079 | -20.6842
12.532 | -34.6083 | -20.6926
12.5382 | -34.6088 | -20.7011
12.5443 | -34.6092 | -20.7094
12.5504 | -34.6096 | -20.7179
12.5565 | -34.6101 | -20.7262
12.5626 | -34.6105 | -20.7347
12.5687 | -34.611 | -20.7432
12.5748 | -34.6114 | -20.7516
12.5809 | -34.6118 | -20.7601
12.587 | -34.6123 | -20.7686
12.5931 | -34.6128 | -20.7767
12.5992 | -34.6132 | -20.7852
12.6053 | -34.6137 | -20.7936
12.6114 | -34.6141 | -20.8019
12.6175 | -34.6146 | -20.8104
12.6236 | -34.615 | -20.8188
12.6297 | -34.6155 | -20.8271
12.6358 | -34.6159 | -20.8354
12.6419 | -34.6163 | -20.8438
12.648 | -34.6168 | -20.8523
12.6541 | -34.6172 | -20.8607
12.6602 | -34.6177 | -20.8689
12.6663 | -34.6181 | -20.8773
12.6724 | -34.6186 | -20.8856
12.6785 | -34.619 | -20.894
12.6847 | -34.6195 | -20.9022
12.6908 | -34.6199 | -20.9107
12.6969 | -34.6204 | -20.919
12.703 | -34.6209 | -20.9276
12.7091 | -34.6213 | -20.936
12.7152 | -34.6217 | -20.9442
12.7213 | -34.6222 | -20.9526
12.7274 | -34.6226 | -20.961
12.7335 | -34.6231 | -20.9692
12.7396 | -34.6235 | -20.9774
12.7457 | -34.624 | -20.9859
12.7518 | -34.6244 | -20.9942
12.7579 | -34.6248 | -21.0027
12.764 | -34.6253 | -21.0112
12.7701 | -34.6257 | -21.0195
12.7762 | -34.6262 | -21.0278
12.7823 | -34.6266 | -21.0362
12.7884 | -34.6271 | -21.0446
12.7945 | -34.6276 | -21.0532
12.8006 | -34.628 | -21.0615
12.8067 | -34.6285 | -21.0698
12.8128 | -34.6289 | -21.0781
12.8189 | -34.6294 | -21.0864
12.8251 | -34.6298 | -21.0948
12.8312 | -34.6302 | -21.1033
12.8373 | -34.6307 | -21.1116
12.8434 | -34.6312 | -21.1201
12.8495 | -34.6316 | -21.1284
12.8556 | -34.632 | -21.1367
12.8617 | -34.6325 | -21.1451
12.8678 | -34.6329 | -21.1536
12.8739 | -34.6334 | -21.1618
12.88 | -34.6338 | -21.1702
12.8861 | -34.6342 | -21.1785
12.8922 | -34.6347 | -21.187
12.8983 | -34.6352 | -21.1953
12.9044 | -34.6356 | -21.2036
12.9105 | -34.6361 | -21.2119
12.9166 | -34.6365 | -21.2203
12.9227 | -34.637 | -21.2286
12.9288 | -34.6374 | -21.237
12.9349 | -34.6379 | -21.2454
12.941 | -34.6384 | -21.2537
12.9471 | -34.6388 | -21.2621
12.9532 | -34.6393 | -21.2704
12.9593 | -34.6397 | -21.2789
12.9654 | -34.6402 | -21.2874
12.9716 | -34.6406 | -21.2958
12.9777 | -34.6411 | -21.3043
12.9838 | -34.6416 | -21.3126
12.9899 | -34.642 | -21.3209
12.996 | -34.6424 | -21.3292
13.0021 | -34.6429 | -21.3376
13.0082 | -34.6434 | -21.3459
13.0143 | -34.6439 | -21.354
13.0204 | -34.6443 | -21.3622
13.0265 | -34.6448 | -21.3705
13.0326 | -34.6452 | -21.3789
13.0387 | -34.6457 | -21.3874
13.0448 | -34.6461 | -21.3957
13.0509 | -34.6466 | -21.4041
13.057 | -34.6471 | -21.4123
13.0631 | -34.6475 | -21.4205
13.0692 | -34.648 | -21.4288
13.0753 | -34.6485 | -21.4373
13.0814 | -34.6489 | -21.4458
13.0875 | -34.6494 | -21.454
13.0936 | -34.6498 | -21.4625
13.0997 | -34.6503 | -21.4707
13.1058 | -34.6507 | -21.479
13.112 | -34.6512 | -21.4873
13.1181 | -34.6517 | -21.4956
13.1242 | -34.6521 | -21.504
13.1303 | -34.6526 | -21.5124
13.1364 | -34.653 | -21.5207
13.1425 | -34.6535 | -21.5291
13.1486 | -34.6539 | -21.5375
13.1547 | -34.6544 | -21.5458
13.1608 | -34.6548 | -21.554
13.1669 | -34.6553 | -21.5622
13.173 | -34.6557 | -21.5706
13.1791 | -34.6562 | -21.5789
13.1852 | -34.6566 | -21.5872
13.1913 | -34.6571 | -21.5955
13.1974 | -34.6575 | -21.6038
13.2035 | -34.6579 | -21.6121
13.2096 | -34.6584 | -21.6203
13.2157 | -34.6589 | -21.6286
13.2218 | -34.6593 | -21.6368
13.2279 | -34.6598 | -21.6452
13.234 | -34.6603 | -21.6535
13.2401 | -34.6607 | -21.6619
13.2462 | -34.6612 | -21.6702
13.2524 | -34.6616 | -21.6787
13.2585 | -34.6621 | -21.6869
13.2646 | -34.6625 | -21.6953
13.2707 | -34.663 | -21.7037
13.2768 | -34.6635 | -21.7121
13.2829 | -34.6639 | -21.7206
13.289 | -34.6644 | -21.7287
13.2951 | -34.6649 | -21.7371
13.3012 | -34.6653 | -21.7454
13.3073 | -34.6658 | -21.7538
13.3134 | -34.6662 | -21.7623
13.3195 | -34.6666 | -21.7706
13.3256 | -34.6671 | -21.7787
13.3317 | -34.6676 | -21.7871
13.3378 | -34.668 | -21.7954
13.3439 | -34.6685 | -21.8038
13.35 | -34.6689 | -21.8123
13.3561 | -34.6694 | -21.8205
13.3622 | -34.6698 | -21.8288
13.3683 | -34.6703 | -21.8371
13.3744 | -34.6707 | -21.8454
13.3805 | -34.6712 | -21.8538
13.3866 | -34.6717 | -21.862
13.3927 | -34.6721 | -21.8705
13.3989 | -34.6726 | -21.879
13.405 | -34.673 | -21.8872
13.4111 | -34.6735 | -21.8954
13.4172 | -34.6739 | -21.9038
13.4233 | -34.6744 | -21.9123
13.4294 | -34.6748 | -21.9208
13.4355 | -34.6753 | -21.9289
13.4416 | -34.6758 | -21.9372
13.4477 | -34.6762 | -21.9456
13.4538 | -34.6767 | -21.954
13.4599 | -34.6772 | -21.9624
13.466 | -34.6776 | -21.9707
13.4721 | -34.6781 | -21.9792
13.4782 | -34.6785 | -21.9876
13.4843 | -34.679 | -21.996
13.4904 | -34.6794 | -22.0043
13.4965 | -34.6799 | -22.0127
13.5026 | -34.6803 | -22.0211
13.5087 | -34.6808 | -22.0293
13.5148 | -34.6812 | -22.0377
13.5209 | -34.6817 | -22.0462
13.527 | -34.6821 | -22.0545
13.5331 | -34.6826 | -22.0628
13.5393 | -34.6831 | -22.0712
13.5454 | -34.6836 | -22.0795
13.5515 | -34.684 | -22.088
13.5576 | -34.6845 | -22.0963
13.5637 | -34.685 | -22.1047
13.5698 | -34.6854 | -22.1131
13.5759 | -34.6859 | -22.1212
13.582 | -34.6864 | -22.1295
13.5881 | -34.6868 | -22.1378
13.5942 | -34.6873 | -22.1462
13.6003 | -34.6878 | -22.1546
13.6064 | -34.6882 | -22.1631
13.6125 | -34.6887 | -22.1715
13.6186 | -34.6892 | -22.1798
13.6247 | -34.6896 | -22.1882
13.6308 | -34.6901 | -22.1967
13.6369 | -34.6905 | -22.2051
13.643 | -34.691 | -22.2136
13.6491 | -34.6915 | -22.222
13.6552 | -34.6919 | -22.2303
13.6613 | -34.6924 | -22.2386
13.6674 | -34.6928 | -22.247
13.6735 | -34.6933 | -22.2554
13.6796 | -34.6937 | -22.2638
13.6858 | -34.6942 | -22.2722
13.6919 | -34.6947 | -22.2806
13.698 | -34.6951 | -22.2891
13.7041 | -34.6956 | -22.2976
13.7102 | -34.6961 | -22.3059
13.7163 | -34.6965 | -22.3143
13.7224 | -34.697 | -22.3227
13.7285 | -34.6975 | -22.3311
13.7346 | -34.6979 | -22.3395
13.7407 | -34.6984 | -22.3479
13.7468 | -34.6988 | -22.3562
13.7529 | -34.6993 | -22.3646
13.759 | -34.6997 | -22.3731
13.7651 | -34.7002 | -22.3815
13.7712 | -34.7007 | -22.3899
13.7773 | -34.7011 | -22.3984
13.7834 | -34.7016 | -22.4068
13.7895 | -34.7021 | -22.4152
13.7956 | -34.7025 | -22.4237
13.8017 | -34.703 | -22.432
13.8078 | -34.7034 | -22.4404
13.8139 | -34.7039 | -22.4487
13.82 | -34.7043 | -22.457
13.8262 | -34.7048 | -22.4656
13.8323 | -34.7053 | -22.474
13.8384 | -34.7057 | -22.4825
13.8445 | -34.7062 | -22.4909
13.8506 | -34.7067 | -22.4993
13.8567 | -34.7071 | -22.5079
13.8628 | -34.7076 | -22.5163
13.8689 | -34.7081 | -22.5248
13.875 | -34.7085 | -22.5331
13.8811 | -34.709 | -22.5415
13.8872 | -34.7094 | -22.5498
13.8933 | -34.7099 | -22.5581
13.8994 | -34.7104 | -22.5666
13.9055 | -34.7108 | -22.5752
13.9116 | -34.7113 | -22.5835
13.9177 | -34.7117 | -22.592
13.9238 | -34.7122 | -22.6004
13.9299 | -34.7127 | -22.6087
13.936 | -34.7131 | -22.6172
13.9421 | -34.7136 | -22.6255
13.9482 | -34.714 | -22.6338
13.9543 | -34.7145 | -22.6423
13.9604 | -34.7149 | -22.6506
13.9665 | -34.7154 | -22.659
13.9727 | -34.7159 | -22.6676
13.9788 | -34.7164 | -22.6759
13.9849 | -34.7168 | -22.6844
13.991 | -34.7173 | -22.6928
13.9971 | -34.7177 | -22.7011
14.0032 | -34.7182 | -22.7094
14.0093 | -34.7186 | -22.7179
14.0154 | -34.7191 | -22.7263
14.0215 | -34.7196 | -22.7346
14.0276 | -34.72 | -22.7429
14.0337 | -34.7205 | -22.7514
14.0398 | -34.721 | -22.7597
14.0459 | -34.7215 | -22.7681
14.052 | -34.7219 | -22.7766
14.0581 | -34.7223 | -22.785
14.0642 | -34.7228 | -22.7934
14.0703 | -34.7233 | -22.8019
14.0764 | -34.7237 | -22.8101
14.0825 | -34.7242 | -22.8186
14.0886 | -34.7247 | -22.827
14.0947 | -34.7251 | -22.8354
14.1008 | -34.7256 | -22.8439
14.1069 | -34.7261 | -22.8522
14.1131 | -34.7265 | -22.8607
14.1192 | -34.727 | -22.869
14.1253 | -34.7274 | -22.8773
14.1314 | -34.7279 | -22.8857
14.1375 | -34.7284 | -22.8941
14.1436 | -34.7289 | -22.9025
14.1497 | -34.7293 | -22.9109
14.1558 | -34.7298 | -22.9193
14.1619 | -34.7303 | -22.9276
14.168 | -34.7307 | -22.936
14.1741 | -34.7312 | -22.9443
14.1802 | -34.7317 | -22.9527
14.1863 | -34.7321 | -22.9611
14.1924 | -34.7326 | -22.9695
14.1985 | -34.7331 | -22.9779
14.2046 | -34.7335 | -22.9864
14.2107 | -34.734 | -22.9947
14.2168 | -34.7345 | -23.0031
14.2229 | -34.7349 | -23.0115
14.229 | -34.7354 | -23.02
14.2351 | -34.7359 | -23.0284
14.2412 | -34.7363 | -23.0368
14.2473 | -34.7368 | -23.0452
14.2534 | -34.7373 | -23.0534
14.2596 | -34.7377 | -23.0618
14.2657 | -34.7382 | -23.0701
14.2718 | -34.7386 | -23.0785
14.2779 | -34.7391 | -23.087
14.284 | -34.7396 | -23.0954
14.2901 | -34.7401 | -23.1038
14.2962 | -34.7405 | -23.1122
14.3023 | -34.741 | -23.1207
14.3084 | -34.7414 | -23.129
14.3145 | -34.7419 | -23.1374
14.3206 | -34.7424 | -23.1459
14.3267 | -34.7429 | -23.1543
14.3328 | -34.7433 | -23.1627
14.3389 | -34.7438 | -23.171
14.345 | -34.7442 | -23.1794
14.3511 | -34.7447 | -23.1877
14.3572 | -34.7452 | -23.1962
14.3633 | -34.7456 | -23.2047
14.3694 | -34.7461 | -23.213
14.3755 | -34.7466 | -23.2215
14.3816 | -34.7471 | -23.2299
14.3877 | -34.7476 | -23.2383
14.3938 | -34.748 | -23.2468
14.4 | -34.7485 | -23.255
14.4061 | -34.749 | -23.2635
14.4122 | -34.7494 | -23.2719
14.4183 | -34.7499 | -23.2803
14.4244 | -34.7503 | -23.2887
14.4305 | -34.7508 | -23.2971
14.4366 | -34.7513 | -23.3055
14.4427 | -34.7518 | -23.314
14.4488 | -34.7523 | -23.3225
14.4549 | -34.7527 | -23.3308
14.461 | -34.7532 | -23.3391
14.4671 | -34.7537 | -23.3476
14.4732 | -34.7541 | -23.356
14.4793 | -34.7546 | -23.3644
14.4854 | -34.7551 | -23.3727
14.4915 | -34.7555 | -23.3809
14.4976 | -34.756 | -23.3891
14.5037 | -34.7565 | -23.3974
14.5098 | -34.757 | -23.4058
14.5159 | -34.7574 | -23.4142
14.522 | -34.7579 | -23.4225
14.5281 | -34.7584 | -23.4308
14.5342 | -34.7588 | -23.4392
14.5403 | -34.7593 | -23.4475
14.5465 | -34.7598 | -23.4559
14.5526 | -34.7602 | -23.4644
14.5587 | -34.7607 | -23.4728
14.5648 | -34.7612 | -23.4811
14.5709 | -34.7617 | -23.4894
14.577 | -34.7621 | -23.4977
14.5831 | -34.7625 | -23.5061
14.5892 | -34.763 | -23.5144
14.5953 | -34.7635 | -23.5227
14.6014 | -34.764 | -23.5311
14.6075 | -34.7644 | -23.5394
14.6136 | -34.7649 | -23.5478
14.6197 | -34.7653 | -23.5562
14.6258 | -34.7658 | -23.5646
14.6319 | -34.7662 | -23.5729
14.638 | -34.7667 | -23.5812
14.6441 | -34.7671 | -23.5896
14.6502 | -34.7676 | -23.5981
14.6563 | -34.7681 | -23.6065
14.6624 | -34.7686 | -23.6148
14.6685 | -34.7691 | -23.623
14.6746 | -34.7695 | -23.6313
14.6807 | -34.77 | -23.6395
14.6869 | -34.7705 | -23.6479
14.693 | -34.771 | -23.6562
14.6991 | -34.7714 | -23.6645
14.7052 | -34.7719 | -23.6727
14.7113 | -34.7724 | -23.6811
14.7174 | -34.7729 | -23.6895
14.7235 | -34.7733 | -23.6977
14.7296 | -34.7738 | -23.706
14.7357 | -34.7743 | -23.7143
14.7418 | -34.7747 | -23.7226
14.7479 | -34.7752 | -23.7309
14.754 | -34.7757 | -23.7392
14.7601 | -34.7762 | -23.7475
14.7662 | -34.7766 | -23.7558
14.7723 | -34.7771 | -23.7638
14.7784 | -34.7776 | -23.7722
14.7845 | -34.778 | -23.7805
14.7906 | -34.7785 | -23.7889
14.7967 | -34.779 | -23.7972
14.8028 | -34.7794 | -23.8054
14.8089 | -34.7799 | -23.8137
14.815 | -34.7804 | -23.8221
14.8211 | -34.7809 | -23.8304
14.8272 | -34.7813 | -23.8388
14.8334 | -34.7818 | -23.8471
14.8395 | -34.7823 | -23.8553
14.8456 | -34.7828 | -23.8635
14.8517 | -34.7833 | -23.8718
14.8578 | -34.7837 | -23.8802
14.8639 | -34.7842 | -23.8885
14.87 | -34.7847 | -23.8968
14.8761 | -34.7852 | -23.9051
14.8822 | -34.7856 | -23.9134
14.8883 | -34.7861 | -23.9217
14.8944 | -34.7866 | -23.93
14.9005 | -34.787 | -23.9383
14.9066 | -34.7875 | -23.9466
14.9127 | -34.788 | -23.9547
14.9188 | -34.7885 | -23.963
14.9249 | -34.7889 | -23.9713
14.931 | -34.7894 | -23.9796
14.9371 | -34.7899 | -23.9878
14.9432 | -34.7904 | -23.9961
14.9493 | -34.7908 | -24.0043
14.9554 | -34.7913 | -24.0125
14.9615 | -34.7918 | -24.0208
14.9676 | -34.7922 | -24.0292
14.9738 | -34.7927 | -24.0375
14.9799 | -34.7932 | -24.0459
14.986 | -34.7937 | -24.054
14.9921 | -34.7941 | -24.0622
14.9982 | -34.7946 | -24.0705
15.0043 | -34.7951 | -24.0786
15.0104 | -34.7955 | -24.087
15.0165 | -34.796 | -24.0954
15.0226 | -34.7965 | -24.1034
15.0287 | -34.797 | -24.1117
15.0348 | -34.7975 | -24.1201
15.0409 | -34.7979 | -24.1284
15.047 | -34.7984 | -24.1368
15.0531 | -34.7989 | -24.1449
15.0592 | -34.7994 | -24.1532
15.0653 | -34.7998 | -24.1616
15.0714 | -34.8003 | -24.1698
15.0775 | -34.8008 | -24.178
15.0836 | -34.8013 | -24.1863
15.0897 | -34.8018 | -24.1942
15.0958 | -34.8023 | -24.2024
15.1019 | -34.8027 | -24.2107
15.108 | -34.8032 | -24.2189
15.1141 | -34.8037 | -24.227
15.1203 | -34.8042 | -24.2354
15.1264 | -34.8047 | -24.2435
15.1325 | -34.8052 | -24.2518
15.1386 | -34.8056 | -24.2599
15.1447 | -34.8061 | -24.2681
15.1508 | -34.8066 | -24.2764
15.1569 | -34.8071 | -24.2846
15.163 | -34.8076 | -24.2926
15.1691 | -34.8081 | -24.3007
15.1752 | -34.8085 | -24.3088
15.1813 | -34.809 | -24.317
15.1874 | -34.8095 | -24.3253
15.1935 | -34.81 | -24.3333
15.1996 | -34.8104 | -24.3414
15.2057 | -34.8109 | -24.3496
15.2118 | -34.8114 | -24.3578
15.2179 | -34.8119 | -24.3661
15.224 | -34.8123 | -24.3742
15.2301 | -34.8128 | -24.3821
15.2362 | -34.8133 | -24.3904
15.2423 | -34.8138 | -24.3987
15.2484 | -34.8143 | -24.4069
15.2545 | -34.8147 | -24.4151
15.2607 | -34.8152 | -24.4233
15.2668 | -34.8157 | -24.4311
15.2729 | -34.8162 | -24.4393
15.279 | -34.8167 | -24.4475
15.2851 | -34.8172 | -24.4556
15.2912 | -34.8177 | -24.4638
15.2973 | -34.8182 | -24.4721
15.3034 | -34.8187 | -24.4802
15.3095 | -34.8192 | -24.4884
15.3156 | -34.8197 | -24.4965
15.3217 | -34.8202 | -24.5046
15.3278 | -34.8206 | -24.5129
15.3339 | -34.8212 | -24.5208
15.34 | -34.8216 | -24.5289
15.3461 | -34.8221 | -24.5369
15.3522 | -34.8226 | -24.545
15.3583 | -34.8231 | -24.5534
15.3644 | -34.8236 | -24.5614
15.3705 | -34.8241 | -24.5692
15.3766 | -34.8245 | -24.5775
15.3827 | -34.825 | -24.5855
15.3888 | -34.8255 | -24.5936
15.3949 | -34.826 | -24.6017
15.401 | -34.8265 | -24.61
15.4072 | -34.827 | -24.6178
15.4133 | -34.8274 | -24.6259
15.4194 | -34.8279 | -24.6341
15.4255 | -34.8284 | -24.6423
15.4316 | -34.8289 | -24.6504
15.4377 | -34.8293 | -24.6586
15.4438 | -34.8299 | -24.6666
15.4499 | -34.8303 | -24.6746
15.456 | -34.8308 | -24.6826
15.4621 | -34.8313 | -24.6907
15.4682 | -34.8318 | -24.699
15.4743 | -34.8323 | -24.7069
15.4804 | -34.8327 | -24.7151
15.4865 | -34.8332 | -24.7231
15.4926 | -34.8337 | -24.7313
15.4987 | -34.8342 | -24.7395
15.5048 | -34.8347 | -24.7476
15.5109 | -34.8352 | -24.7555
15.517 | -34.8357 | -24.7637
15.5231 | -34.8362 | -24.7718
15.5292 | -34.8366 | -24.7798
15.5353 | -34.8371 | -24.788
15.5414 | -34.8376 | -24.796
15.5476 | -34.8381 | -24.8037
15.5537 | -34.8386 | -24.8118
15.5598 | -34.8391 | -24.8199
15.5659 | -34.8396 | -24.828
15.572 | -34.8401 | -24.836
15.5781 | -34.8406 | -24.844
15.5842 | -34.8411 | -24.8518
15.5903 | -34.8416 | -24.8598
15.5964 | -34.8421 | -24.8677
15.6025 | -34.8425 | -24.8757
15.6086 | -34.843 | -24.8837
15.6147 | -34.8435 | -24.8919
15.6208 | -34.844 | -24.8997
15.6269 | -34.8445 | -24.9076
15.633 | -34.845 | -24.9158
15.6391 | -34.8455 | -24.9239
15.6452 | -34.846 | -24.932
15.6513 | -34.8465 | -24.9397
15.6574 | -34.847 | -24.9477
15.6635 | -34.8475 | -24.9557
15.6696 | -34.8479 | -24.9637
15.6757 | -34.8484 | -24.9718
15.6818 | -34.8489 | -24.9798
15.688 | -34.8494 | -24.9876
15.6941 | -34.8499 | -24.9955
15.7002 | -34.8504 | -25.0035
15.7063 | -34.8508 | -25.0115
15.7124 | -34.8513 | -25.0196
15.7185 | -34.8518 | -25.0275
15.7246 | -34.8523 | -25.0355
15.7307 | -34.8528 | -25.0435
15.7368 | -34.8533 | -25.0515
15.7429 | -34.8538 | -25.0594
15.749 | -34.8543 | -25.0674
15.7551 | -34.8548 | -25.0754
15.7612 | -34.8553 | -25.0832
15.7673 | -34.8558 | -25.0912
15.7734 | -34.8563 | -25.0991
15.7795 | -34.8568 | -25.107
15.7856 | -34.8573 | -25.1149
15.7917 | -34.8578 | -25.1225
15.7978 | -34.8583 | -25.1306
15.8039 | -34.8588 | -25.1386
15.81 | -34.8593 | -25.1466
15.8161 | -34.8597 | -25.1545
15.8222 | -34.8602 | -25.1624
15.8283 | -34.8608 | -25.17
15.8345 | -34.8613 | -25.1779
15.8406 | -34.8618 | -25.1857
15.8467 | -34.8623 | -25.1938
15.8528 | -34.8628 | -25.2017
15.8589 | -34.8633 | -25.2097
15.865 | -34.8638 | -25.2173
15.8711 | -34.8643 | -25.2253
15.8772 | -34.8648 | -25.2332
15.8833 | -34.8653 | -25.2413
15.8894 | -34.8658 | -25.2493
15.8955 | -34.8663 | -25.2571
15.9016 | -34.8668 | -25.2648
15.9077 | -34.8672 | -25.2727
15.9138 | -34.8677 | -25.2805
15.9199 | -34.8682 | -25.2884
15.926 | -34.8687 | -25.2963
15.9321 | -34.8692 | -25.304
15.9382 | -34.8697 | -25.3118
15.9443 | -34.8702 | -25.3198
15.9504 | -34.8707 | -25.3278
15.9565 | -34.8712 | -25.3356
15.9626 | -34.8717 | -25.3435
15.9687 | -34.8722 | -25.3511
15.9749 | -34.8727 | -25.3589
15.981 | -34.8731 | -25.3669
15.9871 | -34.8736 | -25.3749
15.9932 | -34.8741 | -25.3828
15.9993 | -34.8746 | -25.3907
16.0054 | -34.8751 | -25.3981
16.0115 | -34.8756 | -25.406
16.0176 | -34.8761 | -25.4139
16.0237 | -34.8766 | -25.4219
16.0298 | -34.8771 | -25.4296
16.0359 | -34.8776 | -25.4374
16.042 | -34.8781 | -25.4451
16.0481 | -34.8786 | -25.453
16.0542 | -34.8791 | -25.461
16.0603 | -34.8796 | -25.4687
16.0664 | -34.88 | -25.4766
16.0725 | -34.8805 | -25.4842
16.0786 | -34.8811 | -25.4921
16.0847 | -34.8815 | -25.5001
16.0908 | -34.882 | -25.5081
16.0969 | -34.8825 | -25.516
16.103 | -34.883 | -25.5239
16.1091 | -34.8835 | -25.5314
16.1152 | -34.884 | -25.5391
16.1214 | -34.8845 | -25.5471
16.1275 | -34.885 | -25.5549
16.1336 | -34.8855 | -25.5625
16.1397 | -34.886 | -25.5703
16.1458 | -34.8865 | -25.5779
16.1519 | -34.887 | -25.5857
16.158 | -34.8875 | -25.5935
16.1641 | -34.8879 | -25.6015
16.1702 | -34.8884 | -25.6094
16.1763 | -34.8889 | -25.6171
16.1824 | -34.8894 | -25.6247
16.1885 | -34.8899 | -25.6325
16.1946 | -34.8904 | -25.6404
16.2007 | -34.8909 | -25.6483
16.2068 | -34.8914 | -25.656
16.2129 | -34.8919 | -25.6639
16.219 | -34.8924 | -25.6713
16.2251 | -34.8929 | -25.6789
16.2312 | -34.8935 | -25.6867
16.2373 | -34.8939 | -25.6945
16.2434 | -34.8944 | -25.7022
16.2495 | -34.895 | -25.7098
16.2556 | -34.8955 | -25.7177
16.2618 | -34.896 | -25.7253
16.2679 | -34.8965 | -25.7331
16.274 | -34.897 | -25.7408
16.2801 | -34.8975 | -25.7486
16.2862 | -34.898 | -25.7563
16.2923 | -34.8985 | -25.7641
16.2984 | -34.899 | -25.7719
16.3045 | -34.8995 | -25.7796
16.3106 | -34.9 | -25.7873
16.3167 | -34.9004 | -25.7951
16.3228 | -34.9009 | -25.8027
16.3289 | -34.9014 | -25.8104
16.335 | -34.9019 | -25.8181
16.3411 | -34.9024 | -25.8259
16.3472 | -34.9029 | -25.8336
16.3533 | -34.9034 | -25.8414
16.3594 | -34.9039 | -25.8489
16.3655 | -34.9044 | -25.8565
16.3716 | -34.9049 | -25.8641
16.3777 | -34.9054 | -25.872
16.3838 | -34.9059 | -25.8799
16.3899 | -34.9064 | -25.8875
16.396 | -34.9069 | -25.8954
16.4021 | -34.9074 | -25.903
16.4083 | -34.9079 | -25.9107
16.4144 | -34.9084 | -25.9185
16.4205 | -34.9089 | -25.9263
16.4266 | -34.9094 | -25.9338
16.4327 | -34.9099 | -25.9415
16.4388 | -34.9104 | -25.9492
16.4449 | -34.9109 | -25.9568
16.451 | -34.9114 | -25.9647
16.4571 | -34.9119 | -25.9726
16.4632 | -34.9124 | -25.9801
16.4693 | -34.9129 | -25.9878
16.4754 | -34.9134 | -25.9955
16.4815 | -34.9139 | -26.0031
16.4876 | -34.9144 | -26.0107
16.4937 | -34.9149 | -26.0184
16.4998 | -34.9154 | -26.0259
16.5059 | -34.9159 | -26.0336
16.512 | -34.9164 | -26.0414
16.5181 | -34.9169 | -26.0493
16.5242 | -34.9174 | -26.0571
16.5303 | -34.9178 | -26.0647
16.5364 | -34.9183 | -26.0724
16.5425 | -34.9188 | -26.08
16.5487 | -34.9193 | -26.0877
16.5548 | -34.9198 | -26.0954
16.5609 | -34.9203 | -26.1031
16.567 | -34.9208 | -26.1105
16.5731 | -34.9214 | -26.1183
16.5792 | -34.9218 | -26.126
16.5853 | -34.9223 | -26.1336
16.5914 | -34.9228 | -26.1414
16.5975 | -34.9233 | -26.149
16.6036 | -34.9238 | -26.1565
16.6097 | -34.9243 | -26.1643
16.6158 | -34.9247 | -26.1719
16.6219 | -34.9252 | -26.1796
16.628 | -34.9257 | -26.1872
16.6341 | -34.9262 | -26.195
16.6402 | -34.9267 | -26.2025
16.6463 | -34.9272 | -26.2102
16.6524 | -34.9277 | -26.2177
16.6585 | -34.9282 | -26.2253
16.6646 | -34.9287 | -26.2329
16.6707 | -34.9292 | -26.2404
16.6768 | -34.9297 | -26.248
16.6829 | -34.9302 | -26.2556
16.689 | -34.9307 | -26.2634
16.6952 | -34.9312 | -26.2711
16.7013 | -34.9317 | -26.2788
16.7074 | -34.9322 | -26.2862
16.7135 | -34.9327 | -26.2938
16.7196 | -34.9332 | -26.3014
16.7257 | -34.9337 | -26.3091
16.7318 | -34.9342 | -26.3167
16.7379 | -34.9347 | -26.3243
16.744 | -34.9352 | -26.3317
16.7501 | -34.9357 | -26.3394
16.7562 | -34.9362 | -26.3471
16.7623 | -34.9367 | -26.3545
16.7684 | -34.9372 | -26.3622
16.7745 | -34.9377 | -26.3698
16.7806 | -34.9382 | -26.3771
16.7867 | -34.9387 | -26.3847
16.7928 | -34.9392 | -26.3923
16.7989 | -34.9396 | -26.3997
16.805 | -34.9401 | -26.4074
16.8111 | -34.9406 | -26.4147
16.8172 | -34.9411 | -26.4224
16.8233 | -34.9417 | -26.4299
16.8294 | -34.9421 | -26.4375
16.8356 | -34.9426 | -26.4451
16.8417 | -34.9431 | -26.4526
16.8478 | -34.9436 | -26.4601
16.8539 | -34.9441 | -26.4676
16.86 | -34.9446 | -26.4754
16.8661 | -34.9451 | -26.483
16.8722 | -34.9456 | -26.4906
16.8783 | -34.9461 | -26.4982
16.8844 | -34.9466 | -26.5055
16.8905 | -34.9471 | -26.5132
16.8966 | -34.9476 | -26.5206
16.9027 | -34.9481 | -26.5282
16.9088 | -34.9486 | -26.5358
16.9149 | -34.9491 | -26.5435
16.921 | -34.9496 | -26.5509
16.9271 | -34.9501 | -26.5587
16.9332 | -34.9505 | -26.5662
16.9393 | -34.951 | -26.5739
16.9454 | -34.9515 | -26.5814
16.9515 | -34.952 | -26.589
16.9576 | -34.9525 | -26.5963
16.9637 | -34.9531 | -26.6037
16.9698 | -34.9536 | -26.611
16.9759 | -34.954 | -26.6184
16.9821 | -34.9545 | -26.626
16.9882 | -34.955 | -26.6333
16.9943 | -34.9555 | -26.6406
17.0004 | -34.956 | -26.6481
17.0065 | -34.9565 | -26.6557
17.0126 | -34.957 | -26.6633
17.0187 | -34.9575 | -26.671
17.0248 | -34.958 | -26.6785
17.0309 | -34.9585 | -26.686
17.037 | -34.959 | -26.6935
17.0431 | -34.9595 | -26.7011
17.0492 | -34.9599 | -26.7086
17.0553 | -34.9604 | -26.7161
17.0614 | -34.9609 | -26.7235
17.0675 | -34.9614 | -26.7309
17.0736 | -34.9619 | -26.7385
17.0797 | -34.9624 | -26.7459
17.0858 | -34.9629 | -26.7534
17.0919 | -34.9634 | -26.761
17.098 | -34.9639 | -26.7683
17.1041 | -34.9643 | -26.7756
17.1102 | -34.9649 | -26.783
17.1163 | -34.9654 | -26.7904
17.1225 | -34.9658 | -26.7978
17.1286 | -34.9663 | -26.805
17.1347 | -34.9668 | -26.8125
17.1408 | -34.9673 | -26.8201
17.1469 | -34.9678 | -26.8276
17.153 | -34.9683 | -26.835
17.1591 | -34.9688 | -26.8425
17.1652 | -34.9693 | -26.8497
17.1713 | -34.9697 | -26.8571
17.1774 | -34.9702 | -26.8646
17.1835 | -34.9707 | -26.8721
17.1896 | -34.9712 | -26.8795
17.1957 | -34.9717 | -26.887
17.2018 | -34.9722 | -26.8941
17.2079 | -34.9727 | -26.9017
17.214 | -34.9732 | -26.9093
17.2201 | -34.9736 | -26.9167
17.2262 | -34.9741 | -26.9241
17.2323 | -34.9746 | -26.9317
17.2384 | -34.9751 | -26.9389
17.2445 | -34.9756 | -26.9463
17.2506 | -34.9761 | -26.9538
17.2567 | -34.9766 | -26.9614
17.2628 | -34.977 | -26.9689
17.269 | -34.9775 | -26.9764
17.2751 | -34.978 | -26.9838
17.2812 | -34.9785 | -26.9913
17.2873 | -34.979 | -26.9989
17.2934 | -34.9795 | -27.0064
17.2995 | -34.98 | -27.0139
17.3056 | -34.9805 | -27.0212
17.3117 | -34.981 | -27.0287
17.3178 | -34.9815 | -27.036
17.3239 | -34.9819 | -27.0435
17.33 | -34.9824 | -27.0508
17.3361 | -34.9829 | -27.0583
17.3422 | -34.9834 | -27.0656
17.3483 | -34.9839 | -27.0731
17.3544 | -34.9844 | -27.0806
17.3605 | -34.9848 | -27.088
17.3666 | -34.9853 | -27.0953
17.3727 | -34.9858 | -27.1027
17.3788 | -34.9863 | -27.1098
17.3849 | -34.9868 | -27.1172
17.391 | -34.9873 | -27.1246
17.3971 | -34.9878 | -27.1321
17.4032 | -34.9882 | -27.1395
17.4094 | -34.9887 | -27.1466
17.4155 | -34.9892 | -27.154
17.4216 | -34.9897 | -27.1614
17.4277 | -34.9902 | -27.1689
17.4338 | -34.9907 | -27.1763
17.4399 | -34.9912 | -27.1838
17.446 | -34.9917 | -27.1911
17.4521 | -34.9922 | -27.1985
17.4582 | -34.9927 | -27.2058
17.4643 | -34.9932 | -27.2131
17.4704 | -34.9936 | -27.2206
17.4765 | -34.9941 | -27.2281
17.4826 | -34.9946 | -27.2354
17.4887 | -34.9951 | -27.2427
17.4948 | -34.9956 | -27.2501
17.5009 | -34.996 | -27.2575
17.507 | -34.9965 | -27.265
17.5131 | -34.997 | -27.2723
17.5192 | -34.9975 | -27.2795
17.5253 | -34.998 | -27.2869
17.5314 | -34.9985 | -27.2942
17.5375 | -34.999 | -27.3016
17.5436 | -34.9994 | -27.3091
17.5497 | -34.9999 | -27.3164
17.5559 | -35.0004 | -27.3236
17.562 | -35.0009 | -27.331
17.5681 | -35.0014 | -27.3384
17.5742 | -35.0019 | -27.3459
17.5803 | -35.0024 | -27.3531
17.5864 | -35.0029 | -27.3605
17.5925 | -35.0034 | -27.368
17.5986 | -35.0039 | -27.3754
17.6047 | -35.0044 | -27.3828
17.6108 | -35.0049 | -27.3902
17.6169 | -35.0053 | -27.3975
17.623 | -35.0058 | -27.4047
17.6291 | -35.0063 | -27.412
17.6352 | -35.0068 | -27.4194
17.6413 | -35.0073 | -27.4268
17.6474 | -35.0078 | -27.4341
17.6535 | -35.0082 | -27.4414
17.6596 | -35.0087 | -27.4485
17.6657 | -35.0092 | -27.4559
17.6718 | -35.0097 | -27.4633
17.6779 | -35.0102 | -27.4705
17.684 | -35.0107 | -27.4779
17.6901 | -35.0112 | -27.4852
17.6963 | -35.0116 | -27.4924
17.7024 | -35.0121 | -27.4997
17.7085 | -35.0126 | -27.5069
17.7146 | -35.0131 | -27.5142
17.7207 | -35.0136 | -27.5214
17.7268 | -35.0141 | -27.5285
17.7329 | -35.0145 | -27.5359
17.739 | -35.015 | -27.5432
17.7451 | -35.0155 | -27.5504
17.7512 | -35.016 | -27.5578
17.7573 | -35.0165 | -27.5652
17.7634 | -35.017 | -27.5723
17.7695 | -35.0175 | -27.5796
17.7756 | -35.018 | -27.5869
17.7817 | -35.0184 | -27.5942
17.7878 | -35.0189 | -27.6015
17.7939 | -35.0194 | -27.609
17.8 | -35.0199 | -27.6161
17.8061 | -35.0204 | -27.6234
17.8122 | -35.0209 | -27.6307
17.8183 | -35.0213 | -27.638
17.8244 | -35.0218 | -27.6454
17.8305 | -35.0223 | -27.6527
17.8366 | -35.0228 | -27.6597
17.8428 | -35.0233 | -27.667
17.8489 | -35.0238 | -27.6743
17.855 | -35.0243 | -27.6816
17.8611 | -35.0247 | -27.6888
17.8672 | -35.0252 | -27.6958
17.8733 | -35.0257 | -27.7032
17.8794 | -35.0262 | -27.7105
17.8855 | -35.0267 | -27.7177
17.8916 | -35.0271 | -27.725
17.8977 | -35.0276 | -27.7325
17.9038 | -35.0281 | -27.7395
17.9099 | -35.0286 | -27.7467
17.916 | -35.0291 | -27.7541
17.9221 | -35.0295 | -27.7613
17.9282 | -35.03 | -27.7686
17.9343 | -35.0305 | -27.776
17.9404 | -35.031 | -27.7832
17.9465 | -35.0315 | -27.7905
17.9526 | -35.0319 | -27.7979
17.9587 | -35.0324 | -27.8051
17.9648 | -35.0329 | -27.8123
17.9709 | -35.0334 | -27.8194
17.977 | -35.0338 | -27.8265
17.9832 | -35.0343 | -27.8337
17.9893 | -35.0348 | -27.841
17.9954 | -35.0353 | -27.8483
18.0015 | -35.0357 | -27.8556
18.0076 | -35.0362 | -27.8627
18.0137 | -35.0367 | -27.8701
18.0198 | -35.0372 | -27.8773
18.0259 | -35.0377 | -27.8846
18.032 | -35.0381 | -27.8918
18.0381 | -35.0386 | -27.8991
18.0442 | -35.0391 | -27.9064
18.0503 | -35.0396 | -27.9137
18.0564 | -35.0401 | -27.9207
18.0625 | -35.0405 | -27.9281
18.0686 | -35.041 | -27.9354
18.0747 | -35.0414 | -27.9429
18.0808 | -35.0419 | -27.95
18.0869 | -35.0424 | -27.9572
18.093 | -35.0429 | -27.9645
18.0991 | -35.0433 | -27.9718
18.1052 | -35.0438 | -27.9791
18.1113 | -35.0443 | -27.9863
18.1174 | -35.0448 | -27.9933
18.1236 | -35.0452 | -28.0007
18.1297 | -35.0457 | -28.0079
18.1358 | -35.0462 | -28.0151
18.1419 | -35.0467 | -28.0223
18.148 | -35.0472 | -28.0295
18.1541 | -35.0477 | -28.0366
18.1602 | -35.0482 | -28.0439
18.1663 | -35.0487 | -28.0511
18.1724 | -35.0491 | -28.0583
18.1785 | -35.0496 | -28.0655
18.1846 | -35.05 | -28.0728
18.1907 | -35.0505 | -28.0801
18.1968 | -35.051 | -28.0873
18.2029 | -35.0515 | -28.0945
18.209 | -35.0519 | -28.1018
18.2151 | -35.0524 | -28.1091
18.2212 | -35.0529 | -28.1162
18.2273 | -35.0534 | -28.1234
18.2334 | -35.0539 | -28.1305
18.2395 | -35.0544 | -28.1378
18.2456 | -35.0548 | -28.145
18.2517 | -35.0553 | -28.1524
18.2578 | -35.0558 | -28.1594
18.2639 | -35.0563 | -28.1667
18.2701 | -35.0567 | -28.1739
18.2762 | -35.0572 | -28.1811
18.2823 | -35.0577 | -28.1882
18.2884 | -35.0582 | -28.1954
18.2945 | -35.0587 | -28.2024
18.3006 | -35.0592 | -28.2097
18.3067 | -35.0597 | -28.2171
18.3128 | -35.0601 | -28.2243
18.3189 | -35.0606 | -28.2315
18.325 | -35.0611 | -28.2384
18.3311 | -35.0616 | -28.2456
18.3372 | -35.0621 | -28.2529
18.3433 | -35.0626 | -28.2601
18.3494 | -35.063 | -28.2672
18.3555 | -35.0635 | -28.2743
18.3616 | -35.064 | -28.2812
18.3677 | -35.0645 | -28.2885
18.3738 | -35.065 | -28.2957
18.3799 | -35.0655 | -28.3027
18.386 | -35.0659 | -28.31
18.3921 | -35.0664 | -28.3173
18.3982 | -35.0668 | -28.3243
18.4043 | -35.0673 | -28.3316
18.4105 | -35.0678 | -28.3388
18.4166 | -35.0683 | -28.346
18.4227 | -35.0687 | -28.3532
18.4288 | -35.0692 | -28.3605
18.4349 | -35.0697 | -28.3674
18.441 | -35.0701 | -28.3746
18.4471 | -35.0706 | -28.3817
18.4532 | -35.0711 | -28.389
18.4593 | -35.0715 | -28.3961
18.4654 | -35.072 | -28.403
18.4715 | -35.0724 | -28.4104
18.4776 | -35.0729 | -28.4176
18.4837 | -35.0733 | -28.4247
18.4898 | -35.0738 | -28.432
18.4959 | -35.0743 | -28.4392
18.502 | -35.0747 | -28.4463
18.5081 | -35.0752 | -28.4535
18.5142 | -35.0756 | -28.4607
18.5203 | -35.0761 | -28.4677
18.5264 | -35.0765 | -28.4748
18.5325 | -35.077 | -28.4821
18.5386 | -35.0775 | -28.4891
18.5447 | -35.0779 | -28.4962
18.5508 | -35.0784 | -28.5033
18.557 | -35.0789 | -28.5105
18.5631 | -35.0793 | -28.5177
18.5692 | -35.0798 | -28.5248
18.5753 | -35.0803 | -28.5317
18.5814 | -35.0808 | -28.5389
18.5875 | -35.0812 | -28.546
18.5936 | -35.0817 | -28.5529
18.5997 | -35.0821 | -28.56
18.6058 | -35.0826 | -28.5671
18.6119 | -35.0831 | -28.5742
18.618 | -35.0835 | -28.5814
18.6241 | -35.084 | -28.5884
18.6302 | -35.0845 | -28.5957
18.6363 | -35.0849 | -28.6029
18.6424 | -35.0854 | -28.6099
18.6485 | -35.0858 | -28.617
18.6546 | -35.0863 | -28.6241
18.6607 | -35.0868 | -28.6311
18.6668 | -35.0872 | -28.6382
18.6729 | -35.0877 | -28.6455
18.679 | -35.0882 | -28.6523
18.6851 | -35.0886 | -28.6593
18.6912 | -35.0891 | -28.6663
18.6974 | -35.0896 | -28.6734
18.7035 | -35.09 | -28.6805
18.7096 | -35.0905 | -28.6878
18.7157 | -35.091 | -28.6949
18.7218 | -35.0915 | -28.702
18.7279 | -35.0919 | -28.7092
18.734 | -35.0924 | -28.7164
18.7401 | -35.0929 | -28.7234
18.7462 | -35.0934 | -28.7302
18.7523 | -35.0939 | -28.7373
18.7584 | -35.0943 | -28.7446
18.7645 | -35.0948 | -28.7516
18.7706 | -35.0953 | -28.7588
18.7767 | -35.0958 | -28.7659
18.7828 | -35.0962 | -28.7728
18.7889 | -35.0967 | -28.7801
18.795 | -35.0971 | -28.7873
18.8011 | -35.0976 | -28.7943
18.8072 | -35.0981 | -28.8014
18.8133 | -35.0985 | -28.8086
18.8194 | -35.099 | -28.8156
18.8255 | -35.0995 | -28.8227
18.8316 | -35.0999 | -28.8297
18.8377 | -35.1004 | -28.8368
18.8439 | -35.1008 | -28.844
18.85 | -35.1013 | -28.8511
18.8561 | -35.1018 | -28.858
18.8622 | -35.1022 | -28.8653
18.8683 | -35.1027 | -28.8723
18.8744 | -35.1032 | -28.8793
18.8805 | -35.1036 | -28.8865
18.8866 | -35.1041 | -28.8935
18.8927 | -35.1046 | -28.9004
18.8988 | -35.1051 | -28.9074
18.9049 | -35.1055 | -28.9145
18.911 | -35.106 | -28.9214
18.9171 | -35.1065 | -28.9285
18.9232 | -35.107 | -28.9354
18.9293 | -35.1074 | -28.9423
18.9354 | -35.1079 | -28.9493
18.9415 | -35.1084 | -28.9564
18.9476 | -35.1088 | -28.9633
18.9537 | -35.1093 | -28.9703
18.9598 | -35.1097 | -28.9772
18.9659 | -35.1102 | -28.9842
18.972 | -35.1107 | -28.9912
18.9781 | -35.1111 | -28.9982
18.9843 | -35.1116 | -29.0053
18.9904 | -35.112 | -29.0123
18.9965 | -35.1125 | -29.019
19.0026 | -35.113 | -29.0261
19.0087 | -35.1135 | -29.0331
19.0148 | -35.114 | -29.0403
19.0209 | -35.1144 | -29.0472
19.027 | -35.1149 | -29.0541
19.0331 | -35.1153 | -29.0609
19.0392 | -35.1158 | -29.0679
19.0453 | -35.1162 | -29.0748
19.0514 | -35.1167 | -29.0819
19.0575 | -35.1171 | -29.0892
19.0636 | -35.1176 | -29.0961
19.0697 | -35.1181 | -29.1031
19.0758 | -35.1185 | -29.1101
19.0819 | -35.119 | -29.117
19.088 | -35.1195 | -29.124
19.0941 | -35.1199 | -29.1312
19.1002 | -35.1204 | -29.1381
19.1063 | -35.1208 | -29.1452
19.1124 | -35.1213 | -29.1521
19.1185 | -35.1218 | -29.1589
19.1246 | -35.1222 | -29.1659
19.1308 | -35.1227 | -29.173
19.1369 | -35.1232 | -29.1798
19.143 | -35.1236 | -29.1868
19.1491 | -35.1241 | -29.1938
19.1552 | -35.1246 | -29.2008
19.1613 | -35.125 | -29.2078
19.1674 | -35.1255 | -29.2146
19.1735 | -35.1259 | -29.2216
19.1796 | -35.1264 | -29.2286
19.1857 | -35.1269 | -29.2357
19.1918 | -35.1273 | -29.2427
19.1979 | -35.1278 | -29.2499
19.204 | -35.1283 | -29.2566
19.2101 | -35.1287 | -29.2638
19.2162 | -35.1292 | -29.2707
19.2223 | -35.1296 | -29.2777
19.2284 | -35.1301 | -29.2848
19.2345 | -35.1305 | -29.2919
19.2406 | -35.131 | -29.2989
19.2467 | -35.1314 | -29.3059
19.2528 | -35.1318 | -29.313
19.2589 | -35.1323 | -29.3199
19.265 | -35.1328 | -29.327
19.2712 | -35.1332 | -29.3339
19.2773 | -35.1337 | -29.3408
19.2834 | -35.1341 | -29.3477
19.2895 | -35.1346 | -29.3546
19.2956 | -35.135 | -29.3616
19.3017 | -35.1355 | -29.3685
19.3078 | -35.1359 | -29.3754
19.3139 | -35.1364 | -29.3822
19.32 | -35.1369 | -29.3892
19.3261 | -35.1373 | -29.3961
19.3322 | -35.1378 | -29.403
19.3383 | -35.1383 | -29.4099
19.3444 | -35.1387 | -29.4168
19.3505 | -35.1392 | -29.4237
19.3566 | -35.1396 | -29.4308
19.3627 | -35.1401 | -29.4376
19.3688 | -35.1405 | -29.4445
19.3749 | -35.141 | -29.4515
19.381 | -35.1414 | -29.4581
19.3871 | -35.1419 | -29.465
19.3932 | -35.1423 | -29.472
19.3993 | -35.1428 | -29.4789
19.4054 | -35.1432 | -29.4859
19.4115 | -35.1437 | -29.4929
19.4177 | -35.1441 | -29.4995
19.4238 | -35.1446 | -29.5064
19.4299 | -35.1451 | -29.5132
19.436 | -35.1455 | -29.5203
19.4421 | -35.146 | -29.5271
19.4482 | -35.1465 | -29.5341
19.4543 | -35.1469 | -29.5407
19.4604 | -35.1473 | -29.5475
19.4665 | -35.1478 | -29.5543
19.4726 | -35.1483 | -29.5612
19.4787 | -35.1487 | -29.5682
19.4848 | -35.1492 | -29.5752
19.4909 | -35.1496 | -29.5821
19.497 | -35.1501 | -29.589
19.5031 | -35.1505 | -29.5959
19.5092 | -35.151 | -29.6028
19.5153 | -35.1514 | -29.6097
19.5214 | -35.1519 | -29.6164
19.5275 | -35.1523 | -29.6233
19.5336 | -35.1528 | -29.6302
19.5397 | -35.1532 | -29.6371
19.5458 | -35.1537 | -29.644
19.5519 | -35.1541 | -29.6507
19.5581 | -35.1546 | -29.6575
19.5642 | -35.155 | -29.6644
19.5703 | -35.1554 | -29.6714
19.5764 | -35.1559 | -29.6783
19.5825 | -35.1564 | -29.6852
19.5886 | -35.1568 | -29.6921
19.5947 | -35.1573 | -29.6988
19.6008 | -35.1577 | -29.7056
19.6069 | -35.1582 | -29.7126
19.613 | -35.1586 | -29.7198
19.6191 | -35.1591 | -29.7267
19.6252 | -35.1595 | -29.7336
19.6313 | -35.16 | -29.7402
19.6374 | -35.1604 | -29.7471
19.6435 | -35.1608 | -29.7538
19.6496 | -35.1613 | -29.7606
19.6557 | -35.1617 | -29.7675
19.6618 | -35.1622 | -29.7742
19.6679 | -35.1626 | -29.781
19.674 | -35.1631 | -29.7879
19.6801 | -35.1635 | -29.7948
19.6862 | -35.164 | -29.8017
19.6923 | -35.1645 | -29.8084
19.6984 | -35.1649 | -29.815
19.7046 | -35.1654 | -29.8219
19.7107 | -35.1659 | -29.8286
19.7168 | -35.1663 | -29.8356
19.7229 | -35.1667 | -29.8424
19.729 | -35.1672 | -29.849
19.7351 | -35.1676 | -29.8557
19.7412 | -35.1681 | -29.8627
19.7473 | -35.1686 | -29.8695
19.7534 | -35.169 | -29.8763
19.7595 | -35.1694 | -29.8831
19.7656 | -35.1699 | -29.8899
19.7717 | -35.1703 | -29.8965
19.7778 | -35.1707 | -29.9033
19.7839 | -35.1712 | -29.9101
19.79 | -35.1717 | -29.9168
19.7961 | -35.1721 | -29.9235
19.8022 | -35.1726 | -29.9303
19.8083 | -35.1731 | -29.9371
19.8144 | -35.1735 | -29.9438
19.8205 | -35.1739 | -29.9506
19.8266 | -35.1744 | -29.9575
19.8327 | -35.1749 | -29.9641
19.8388 | -35.1754 | -29.9708
19.845 | -35.1758 | -29.9775
19.8511 | -35.1763 | -29.9843
19.8572 | -35.1767 | -29.9911
19.8633 | -35.1771 | -29.9977
19.8694 | -35.1776 | -30.0046
19.8755 | -35.178 | -30.011
19.8816 | -35.1785 | -30.0179
19.8877 | -35.1789 | -30.0249
19.8938 | -35.1794 | -30.0317
19.8999 | -35.1798 | -30.0385
19.906 | -35.1803 | -30.0453
19.9121 | -35.1807 | -30.0518
19.9182 | -35.1811 | -30.0586
19.9243 | -35.1816 | -30.0653
19.9304 | -35.182 | -30.0719
19.9365 | -35.1824 | -30.0788
19.9426 | -35.1829 | -30.0852
19.9487 | -35.1833 | -30.0919
19.9548 | -35.1838 | -30.0989
19.9609 | -35.1842 | -30.1055
19.967 | -35.1847 | -30.1121
19.9731 | -35.1852 | -30.1187
19.9792 | -35.1856 | -30.1256
19.9853 | -35.1861 | -30.1324
19.9915 | -35.1865 | -30.1392
19.9976 | -35.1869 | -30.146
20.0037 | -35.1874 | -30.1528
20.0098 | -35.1878 | -30.1597
20.0159 | -35.1883 | -30.1664
20.022 | -35.1887 | -30.173
20.0281 | -35.1891 | -30.1798
20.0342 | -35.1896 | -30.1865
20.0403 | -35.19 | -30.1933
20.0464 | -35.1905 | -30.2001
20.0525 | -35.1909 | -30.2065
20.0586 | -35.1914 | -30.2131
20.0647 | -35.1918 | -30.2198
20.0708 | -35.1923 | -30.2268
20.0769 | -35.1927 | -30.2337
20.083 | -35.1932 | -30.24
20.0891 | -35.1937 | -30.2468
20.0952 | -35.1941 | -30.2536
20.1013 | -35.1945 | -30.2604
20.1074 | -35.195 | -30.2671
20.1135 | -35.1954 | -30.2737
20.1196 | -35.1959 | -30.2801
20.1257 | -35.1963 | -30.2869
20.1319 | -35.1967 | -30.2937
20.138 | -35.1971 | -30.3005
20.1441 | -35.1976 | -30.3073
20.1502 | -35.198 | -30.3142
20.1563 | -35.1985 | -30.3206
20.1624 | -35.1989 | -30.3272
20.1685 | -35.1993 | -30.3342
20.1746 | -35.1998 | -30.3408
20.1807 | -35.2003 | -30.3475
20.1868 | -35.2007 | -30.3542
20.1929 | -35.2011 | -30.3608
20.199 | -35.2016 | -30.3675
20.2051 | -35.202 | -30.3743
20.2112 | -35.2025 | -30.381
20.2173 | -35.2029 | -30.3874
20.2234 | -35.2033 | -30.3942
20.2295 | -35.2038 | -30.4007
20.2356 | -35.2042 | -30.4075
20.2417 | -35.2047 | -30.4144
20.2478 | -35.2051 | -30.4209
20.2539 | -35.2056 | -30.4275
20.26 | -35.206 | -30.4337
20.2661 | -35.2065 | -30.4404
20.2723 | -35.207 | -30.4472
20.2784 | -35.2074 | -30.4539
20.2845 | -35.2079 | -30.4604
20.2906 | -35.2083 | -30.4673
20.2967 | -35.2088 | -30.4736
20.3028 | -35.2092 | -30.4804
20.3089 | -35.2096 | -30.4871
20.315 | -35.2101 | -30.4935
20.3211 | -35.2105 | -30.5002
20.3272 | -35.211 | -30.5069
20.3333 | -35.2114 | -30.5132
20.3394 | -35.2119 | -30.5199
20.3455 | -35.2123 | -30.5264
20.3516 | -35.2127 | -30.5332
20.3577 | -35.2131 | -30.5397
20.3638 | -35.2136 | -30.5464
20.3699 | -35.214 | -30.553
20.376 | -35.2145 | -30.5599
20.3821 | -35.2149 | -30.5666
20.3882 | -35.2154 | -30.5733
20.3943 | -35.2158 | -30.58
20.4004 | -35.2162 | -30.5862
20.4065 | -35.2167 | -30.5929
20.4126 | -35.2171 | -30.5996
20.4188 | -35.2176 | -30.6061
20.4249 | -35.218 | -30.6128
20.431 | -35.2184 | -30.6194
20.4371 | -35.2189 | -30.6257
20.4432 | -35.2193 | -30.6322
20.4493 | -35.2198 | -30.6389
20.4554 | -35.2202 | -30.6456
20.4615 | -35.2207 | -30.6523
20.4676 | -35.2211 | -30.659
20.4737 | -35.2216 | -30.6657
20.4798 | -35.2221 | -30.6724
20.4859 | -35.2225 | -30.679
20.492 | -35.2229 | -30.6855
20.4981 | -35.2234 | -30.6921
20.5042 | -35.2238 | -30.6987
20.5103 | -35.2242 | -30.705
20.5164 | -35.2247 | -30.7116
20.5225 | -35.2251 | -30.7182
20.5286 | -35.2255 | -30.7248
20.5347 | -35.226 | -30.7313
20.5408 | -35.2264 | -30.7377
20.5469 | -35.2269 | -30.7444
20.553 | -35.2273 | -30.7511
20.5592 | -35.2277 | -30.7578
20.5653 | -35.2282 | -30.7644
20.5714 | -35.2286 | -30.7711
20.5775 | -35.2291 | -30.7777
20.5836 | -35.2295 | -30.7844
20.5897 | -35.2299 | -30.7909
20.5958 | -35.2304 | -30.7977
20.6019 | -35.2308 | -30.8043
20.608 | -35.2313 | -30.8111
20.6141 | -35.2317 | -30.8175
20.6202 | -35.2321 | -30.8241
20.6263 | -35.2326 | -30.8305
20.6324 | -35.233 | -30.837
20.6385 | -35.2334 | -30.8435
20.6446 | -35.2339 | -30.8501
20.6507 | -35.2343 | -30.8564
20.6568 | -35.2347 | -30.863
20.6629 | -35.2351 | -30.8696
20.669 | -35.2356 | -30.8761
20.6751 | -35.236 | -30.8826
20.6812 | -35.2365 | -30.8888
20.6873 | -35.2369 | -30.8953
20.6934 | -35.2373 | -30.9019
20.6995 | -35.2378 | -30.9084
20.7057 | -35.2382 | -30.915
20.7118 | -35.2387 | -30.9216
20.7179 | -35.2391 | -30.9279
20.724 | -35.2395 | -30.9343
20.7301 | -35.2399 | -30.9408
20.7362 | -35.2404 | -30.9472
20.7423 | -35.2408 | -30.9536
20.7484 | -35.2413 | -30.9602
20.7545 | -35.2417 | -30.9664
20.7606 | -35.2422 | -30.9731
20.7667 | -35.2426 | -30.9794
20.7728 | -35.2431 | -30.9858
20.7789 | -35.2435 | -30.9924
20.785 | -35.2439 | -30.9991
20.7911 | -35.2444 | -31.0054
20.7972 | -35.2448 | -31.0119
20.8033 | -35.2453 | -31.0185
20.8094 | -35.2457 | -31.025
20.8155 | -35.2462 | -31.0314
20.8216 | -35.2465 | -31.0377
20.8277 | -35.247 | -31.0438
20.8338 | -35.2474 | -31.0502
20.8399 | -35.2478 | -31.0569
20.8461 | -35.2483 | -31.0634
20.8522 | -35.2487 | -31.07
20.8583 | -35.2492 | -31.0763
20.8644 | -35.2496 | -31.0829
20.8705 | -35.2501 | -31.0895
20.8766 | -35.2505 | -31.096
20.8827 | -35.2509 | -31.1024
20.8888 | -35.2513 | -31.109
20.8949 | -35.2518 | -31.1154
20.901 | -35.2522 | -31.122
20.9071 | -35.2526 | -31.1284
20.9132 | -35.2531 | -31.135
20.9193 | -35.2535 | -31.1416
20.9254 | -35.254 | -31.1481
20.9315 | -35.2544 | -31.1545
20.9376 | -35.2549 | -31.161
20.9437 | -35.2553 | -31.1675
20.9498 | -35.2557 | -31.1739
20.9559 | -35.2562 | -31.1804
20.962 | -35.2566 | -31.1868
20.9681 | -35.257 | -31.193
20.9742 | -35.2575 | -31.1995
20.9803 | -35.2579 | -31.206
20.9864 | -35.2583 | -31.2124
20.9926 | -35.2588 | -31.2191
20.9987 | -35.2592 | -31.2254
21.0048 | -35.2596 | -31.2319
21.0109 | -35.2601 | -31.2385
21.017 | -35.2605 | -31.2449
21.0231 | -35.261 | -31.2513
21.0292 | -35.2614 | -31.2578
21.0353 | -35.2619 | -31.2639
21.0414 | -35.2623 | -31.2704
21.0475 | -35.2627 | -31.2766
21.0536 | -35.2632 | -31.2831
21.0597 | -35.2636 | -31.2895
21.0658 | -35.2641 | -31.2959
21.0719 | -35.2645 | -31.3021
21.078 | -35.2649 | -31.3087
21.0841 | -35.2653 | -31.315
21.0902 | -35.2657 | -31.3216
21.0963 | -35.2662 | -31.328
21.1024 | -35.2666 | -31.3346
21.1085 | -35.2671 | -31.3407
21.1146 | -35.2675 | -31.3471
21.1207 | -35.268 | -31.3535
21.1268 | -35.2684 | -31.36
21.133 | -35.2688 | -31.3666
21.1391 | -35.2693 | -31.3728
21.1452 | -35.2697 | -31.3791
21.1513 | -35.2701 | -31.3854
21.1574 | -35.2706 | -31.3918
21.1635 | -35.271 | -31.3983
21.1696 | -35.2715 | -31.4048
21.1757 | -35.2719 | -31.411
21.1818 | -35.2724 | -31.4174
21.1879 | -35.2728 | -31.4239
21.194 | -35.2732 | -31.4304
21.2001 | -35.2737 | -31.4369
21.2062 | -35.2741 | -31.4434
21.2123 | -35.2745 | -31.4492
21.2184 | -35.2749 | -31.4556
21.2245 | -35.2754 | -31.462
21.2306 | -35.2758 | -31.4684
21.2367 | -35.2762 | -31.4749
21.2428 | -35.2767 | -31.4811
21.2489 | -35.2771 | -31.4872
21.255 | -35.2775 | -31.4938
21.2611 | -35.2779 | -31.5001
21.2672 | -35.2784 | -31.5065
21.2733 | -35.2788 | -31.5129
21.2795 | -35.2792 | -31.5189
21.2856 | -35.2797 | -31.5252
21.2917 | -35.2801 | -31.5316
21.2978 | -35.2806 | -31.5378
21.3039 | -35.281 | -31.5443
21.31 | -35.2814 | -31.5508
21.3161 | -35.2819 | -31.5571
21.3222 | -35.2823 | -31.5637
21.3283 | -35.2828 | -31.5699
21.3344 | -35.2832 | -31.5762
21.3405 | -35.2836 | -31.5826
21.3466 | -35.284 | -31.5891
21.3527 | -35.2845 | -31.5953
21.3588 | -35.285 | -31.6015
21.3649 | -35.2854 | -31.6078
21.371 | -35.2859 | -31.6142
21.3771 | -35.2863 | -31.6206
21.3832 | -35.2867 | -31.6269
21.3893 | -35.2872 | -31.633
21.3954 | -35.2876 | -31.6395
21.4015 | -35.288 | -31.6457
21.4076 | -35.2885 | -31.6522
21.4137 | -35.2889 | -31.6587
21.4199 | -35.2893 | -31.6651
21.426 | -35.2898 | -31.6713
21.4321 | -35.2902 | -31.6778
21.4382 | -35.2907 | -31.684
21.4443 | -35.2911 | -31.6904
21.4504 | -35.2915 | -31.6966
21.4565 | -35.2919 | -31.7026
21.4626 | -35.2923 | -31.7091
21.4687 | -35.2928 | -31.7154
21.4748 | -35.2932 | -31.7215
21.4809 | -35.2936 | -31.7278
21.487 | -35.294 | -31.7339
21.4931 | -35.2945 | -31.7399
21.4992 | -35.2949 | -31.7462
21.5053 | -35.2953 | -31.7524
21.5114 | -35.2958 | -31.7588
21.5175 | -35.2962 | -31.7651
21.5236 | -35.2966 | -31.7714
21.5297 | -35.2971 | -31.7777
21.5358 | -35.2975 | -31.7841
21.5419 | -35.2979 | -31.7906
21.548 | -35.2983 | -31.7971
21.5541 | -35.2987 | -31.8035
21.5602 | -35.2992 | -31.81
21.5664 | -35.2996 | -31.8161
21.5725 | -35.3 | -31.8223
21.5786 | -35.3004 | -31.8287
21.5847 | -35.3009 | -31.8352
21.5908 | -35.3013 | -31.8414
21.5969 | -35.3017 | -31.8476
21.603 | -35.3022 | -31.854
21.6091 | -35.3026 | -31.8603
21.6152 | -35.303 | -31.8668
21.6213 | -35.3035 | -31.8732
21.6274 | -35.3039 | -31.8794
21.6335 | -35.3043 | -31.8855
21.6396 | -35.3047 | -31.8919
21.6457 | -35.3051 | -31.8983
21.6518 | -35.3056 | -31.9045
21.6579 | -35.306 | -31.9109
21.664 | -35.3065 | -31.9172
21.6701 | -35.3069 | -31.9232
21.6762 | -35.3073 | -31.9295
21.6823 | -35.3077 | -31.9359
21.6884 | -35.3082 | -31.9423
21.6945 | -35.3086 | -31.9486
21.7006 | -35.309 | -31.9548
21.7068 | -35.3094 | -31.961
21.7129 | -35.3099 | -31.9673
21.719 | -35.3103 | -31.9736
21.7251 | -35.3107 | -31.9799
21.7312 | -35.3111 | -31.9862
21.7373 | -35.3116 | -31.9922
21.7434 | -35.312 | -31.9986
21.7495 | -35.3125 | -32.0049
21.7556 | -35.3129 | -32.0111
21.7617 | -35.3133 | -32.0174
21.7678 | -35.3137 | -32.0237
21.7739 | -35.3141 | -32.0299
21.78 | -35.3146 | -32.0363
21.7861 | -35.315 | -32.0428
21.7922 | -35.3154 | -32.0489
21.7983 | -35.3158 | -32.0554
21.8044 | -35.3163 | -32.0617
21.8105 | -35.3167 | -32.0678
21.8166 | -35.3172 | -32.0741
21.8227 | -35.3176 | -32.0803
21.8288 | -35.3181 | -32.0868
21.8349 | -35.3185 | -32.0931
21.841 | -35.3189 | -32.0994
21.8471 | -35.3194 | -32.1054
21.8533 | -35.3198 | -32.1115
21.8594 | -35.3202 | -32.1179
21.8655 | -35.3207 | -32.124
21.8716 | -35.3211 | -32.1303
21.8777 | -35.3215 | -32.1365
21.8838 | -35.3219 | -32.1428
21.8899 | -35.3223 | -32.149
21.896 | -35.3228 | -32.1551
21.9021 | -35.3232 | -32.1615
21.9082 | -35.3236 | -32.1679
21.9143 | -35.3241 | -32.1741
21.9204 | -35.3245 | -32.1802
21.9265 | -35.3249 | -32.1865
21.9326 | -35.3254 | -32.1929
21.9387 | -35.3258 | -32.199
21.9448 | -35.3262 | -32.2054
21.9509 | -35.3266 | -32.2115
21.957 | -35.3271 | -32.2177
21.9631 | -35.3275 | -32.224
21.9692 | -35.3279 | -32.2301
21.9753 | -35.3283 | -32.2364
21.9814 | -35.3288 | -32.2426
21.9875 | -35.3292 | -32.2488
21.9937 | -35.3296 | -32.2551
21.9998 | -35.3301 | -32.2615
22.0059 | -35.3305 | -32.2678
22.012 | -35.3309 | -32.2741
22.0181 | -35.3314 | -32.28
22.0242 | -35.3318 | -32.2861
22.0303 | -35.3323 | -32.2925
22.0364 | -35.3327 | -32.2986
22.0425 | -35.3331 | -32.3048
22.0486 | -35.3335 | -32.3109
22.0547 | -35.334 | -32.317
22.0608 | -35.3344 | -32.3233
22.0669 | -35.3349 | -32.3295
22.073 | -35.3353 | -32.3356
22.0791 | -35.3357 | -32.3419
22.0852 | -35.3362 | -32.348
22.0913 | -35.3366 | -32.3541
22.0974 | -35.337 | -32.3604
22.1035 | -35.3374 | -32.3666
22.1096 | -35.3378 | -32.3728
22.1157 | -35.3382 | -32.3791
22.1218 | -35.3387 | -32.3853
22.1279 | -35.3391 | -32.3914
22.134 | -35.3395 | -32.3976
22.1402 | -35.3399 | -32.4036
22.1463 | -35.3404 | -32.4098
22.1524 | -35.3408 | -32.4162
22.1585 | -35.3412 | -32.4224
22.1646 | -35.3416 | -32.4283
22.1707 | -35.342 | -32.4346
22.1768 | -35.3425 | -32.4409
22.1829 | -35.3429 | -32.4472
22.189 | -35.3433 | -32.4535
22.1951 | -35.3438 | -32.4597
22.2012 | -35.3442 | -32.466
22.2073 | -35.3447 | -32.4724
22.2134 | -35.3451 | -32.4786
22.2195 | -35.3455 | -32.4848
22.2256 | -35.346 | -32.4912
22.2317 | -35.3464 | -32.4972
22.2378 | -35.3468 | -32.5035
22.2439 | -35.3472 | -32.5097
22.25 | -35.3476 | -32.5159
22.2561 | -35.3481 | -32.5221
22.2622 | -35.3485 | -32.5282
22.2683 | -35.3489 | -32.5344
22.2744 | -35.3493 | -32.5406
22.2806 | -35.3498 | -32.5467
22.2867 | -35.3502 | -32.5528
22.2928 | -35.3506 | -32.559
22.2989 | -35.3511 | -32.5652
22.305 | -35.3515 | -32.5712
22.3111 | -35.3519 | -32.5775
22.3172 | -35.3524 | -32.5837
22.3233 | -35.3528 | -32.5899
22.3294 | -35.3532 | -32.5962
22.3355 | -35.3536 | -32.6024
22.3416 | -35.3541 | -32.6085
22.3477 | -35.3545 | -32.6147
22.3538 | -35.3549 | -32.6208
22.3599 | -35.3553 | -32.6269
22.366 | -35.3558 | -32.6329
22.3721 | -35.3562 | -32.6389
22.3782 | -35.3566 | -32.6451
22.3843 | -35.357 | -32.6514
22.3904 | -35.3574 | -32.6574
22.3965 | -35.3579 | -32.6634
22.4026 | -35.3583 | -32.6696
22.4087 | -35.3587 | -32.6754
22.4148 | -35.3591 | -32.6815
22.4209 | -35.3596 | -32.6877
22.4271 | -35.36 | -32.6938
22.4332 | -35.3604 | -32.6999
22.4393 | -35.3609 | -32.7061
22.4454 | -35.3613 | -32.7121
22.4515 | -35.3617 | -32.7181
22.4576 | -35.3622 | -32.7242
22.4637 | -35.3626 | -32.7306
22.4698 | -35.3631 | -32.7368
22.4759 | -35.3635 | -32.7427
22.482 | -35.3639 | -32.7489
22.4881 | -35.3643 | -32.7552
22.4942 | -35.3647 | -32.7614
22.5003 | -35.3651 | -32.7677
22.5064 | -35.3656 | -32.774
22.5125 | -35.366 | -32.78
22.5186 | -35.3664 | -32.7861
22.5247 | -35.3668 | -32.7924
22.5308 | -35.3673 | -32.7987
22.5369 | -35.3677 | -32.805
22.543 | -35.3681 | -32.8111
22.5491 | -35.3686 | -32.8171
22.5552 | -35.369 | -32.8232
22.5613 | -35.3694 | -32.8293
22.5675 | -35.3699 | -32.8353
22.5736 | -35.3703 | -32.8416
22.5797 | -35.3707 | -32.8479
22.5858 | -35.3711 | -32.8538
22.5919 | -35.3716 | -32.8599
22.598 | -35.372 | -32.8661
22.6041 | -35.3724 | -32.8722
22.6102 | -35.3729 | -32.8786
22.6163 | -35.3733 | -32.8847
22.6224 | -35.3738 | -32.8909
22.6285 | -35.3742 | -32.8971
22.6346 | -35.3746 | -32.9034
22.6407 | -35.375 | -32.9094
22.6468 | -35.3755 | -32.9156
22.6529 | -35.3759 | -32.9215
22.659 | -35.3763 | -32.9274
22.6651 | -35.3768 | -32.9335
22.6712 | -35.3772 | -32.9395
22.6773 | -35.3776 | -32.9456
22.6834 | -35.378 | -32.9518
22.6895 | -35.3784 | -32.9577
22.6956 | -35.3788 | -32.9638
22.7017 | -35.3793 | -32.9701
22.7079 | -35.3797 | -32.976
22.714 | -35.3801 | -32.9822
22.7201 | -35.3805 | -32.9882
22.7262 | -35.3809 | -32.9942
22.7323 | -35.3814 | -33.0003
22.7384 | -35.3818 | -33.0065
22.7445 | -35.3822 | -33.0126
22.7506 | -35.3826 | -33.0187
22.7567 | -35.3831 | -33.0248
22.7628 | -35.3835 | -33.0307
22.7689 | -35.3839 | -33.0368
22.775 | -35.3844 | -33.043
22.7811 | -35.3848 | -33.0492
22.7872 | -35.3852 | -33.0553
22.7933 | -35.3856 | -33.0613
22.7994 | -35.3861 | -33.0674
22.8055 | -35.3865 | -33.0734
22.8116 | -35.387 | -33.0794
22.8177 | -35.3874 | -33.0857
22.8238 | -35.3878 | -33.0917
22.8299 | -35.3882 | -33.0978
22.836 | -35.3886 | -33.1039
22.8421 | -35.3891 | -33.1102
22.8482 | -35.3895 | -33.1162
22.8544 | -35.3899 | -33.1224
22.8605 | -35.3903 | -33.1286
22.8666 | -35.3907 | -33.1346
22.8727 | -35.3912 | -33.1406
22.8788 | -35.3916 | -33.1466
22.8849 | -35.392 | -33.1527
22.891 | -35.3924 | -33.1589
22.8971 | -35.3929 | -33.165
22.9032 | -35.3933 | -33.171
22.9093 | -35.3937 | -33.1773
22.9154 | -35.3941 | -33.1832
22.9215 | -35.3945 | -33.1894
22.9276 | -35.395 | -33.1953
22.9337 | -35.3954 | -33.2011
22.9398 | -35.3958 | -33.2072
22.9459 | -35.3962 | -33.2132
22.952 | -35.3967 | -33.2192
22.9581 | -35.3971 | -33.2253
22.9642 | -35.3975 | -33.2314
22.9703 | -35.3979 | -33.2375
22.9764 | -35.3983 | -33.2436
22.9825 | -35.3988 | -33.2496
22.9886 | -35.3992 | -33.2556
22.9948 | -35.3996 | -33.2617
23.0009 | -35.4 | -33.2678
23.007 | -35.4004 | -33.2738
23.0131 | -35.4008 | -33.2798
23.0192 | -35.4013 | -33.2858
23.0253 | -35.4017 | -33.2918
23.0314 | -35.4021 | -33.2979
23.0375 | -35.4025 | -33.3042
23.0436 | -35.403 | -33.3101
23.0497 | -35.4034 | -33.3163
23.0558 | -35.4038 | -33.3223
23.0619 | -35.4042 | -33.3283
23.068 | -35.4046 | -33.3345
23.0741 | -35.4051 | -33.3403
23.0802 | -35.4055 | -33.3464
23.0863 | -35.4059 | -33.3525
23.0924 | -35.4063 | -33.3587
23.0985 | -35.4068 | -33.3648
23.1046 | -35.4072 | -33.3708
23.1107 | -35.4076 | -33.3769
23.1168 | -35.408 | -33.383
23.1229 | -35.4085 | -33.3891
23.129 | -35.4089 | -33.3953
23.1351 | -35.4093 | -33.4013
23.1413 | -35.4097 | -33.4075
23.1474 | -35.4101 | -33.4135
23.1535 | -35.4105 | -33.4195
23.1596 | -35.411 | -33.4256
23.1657 | -35.4114 | -33.4318
23.1718 | -35.4119 | -33.4379
23.1779 | -35.4123 | -33.444
23.184 | -35.4127 | -33.4499
23.1901 | -35.4131 | -33.4559
23.1962 | -35.4136 | -33.4619
23.2023 | -35.414 | -33.4681
23.2084 | -35.4144 | -33.4742
23.2145 | -35.4148 | -33.4801
23.2206 | -35.4152 | -33.4861
23.2267 | -35.4156 | -33.4922
23.2328 | -35.416 | -33.4982
23.2389 | -35.4165 | -33.5042
23.245 | -35.4169 | -33.5102
23.2511 | -35.4173 | -33.5162
23.2572 | -35.4177 | -33.5223
23.2633 | -35.4181 | -33.5284
23.2694 | -35.4185 | -33.5345
23.2755 | -35.419 | -33.5406
23.2817 | -35.4194 | -33.5466
23.2878 | -35.4198 | -33.5526
23.2939 | -35.4202 | -33.5587
23.3 | -35.4206 | -33.5648
23.3061 | -35.4211 | -33.5709
23.3122 | -35.4215 | -33.577
23.3183 | -35.4219 | -33.583
23.3244 | -35.4223 | -33.589
23.3305 | -35.4228 | -33.5952
23.3366 | -35.4232 | -33.6013
23.3427 | -35.4236 | -33.6072
23.3488 | -35.424 | -33.6132
23.3549 | -35.4245 | -33.6195
23.361 | -35.4249 | -33.6255
23.3671 | -35.4253 | -33.6316
23.3732 | -35.4257 | -33.6376
23.3793 | -35.4261 | -33.6436
23.3854 | -35.4265 | -33.6498
23.3915 | -35.427 | -33.6557
23.3976 | -35.4274 | -33.6617
23.4037 | -35.4278 | -33.6677
23.4098 | -35.4282 | -33.6737
23.4159 | -35.4287 | -33.6798
23.422 | -35.4291 | -33.6859
23.4282 | -35.4295 | -33.6921
23.4343 | -35.43 | -33.6982
23.4404 | -35.4304 | -33.7042
23.4465 | -35.4308 | -33.7103
23.4526 | -35.4312 | -33.7164
23.4587 | -35.4316 | -33.7225
23.4648 | -35.432 | -33.7283
23.4709 | -35.4324 | -33.7344
23.477 | -35.4329 | -33.7405
23.4831 | -35.4333 | -33.7464
23.4892 | -35.4337 | -33.7524
23.4953 | -35.4342 | -33.7584
23.5014 | -35.4346 | -33.7644
23.5075 | -35.435 | -33.7705
23.5136 | -35.4354 | -33.7765
23.5197 | -35.4358 | -33.7824
23.5258 | -35.4362 | -33.7884
23.5319 | -35.4366 | -33.7943
23.538 | -35.437 | -33.8003
23.5441 | -35.4375 | -33.8064
23.5502 | -35.4379 | -33.8124
23.5563 | -35.4383 | -33.8185
23.5624 | -35.4387 | -33.8246
23.5686 | -35.4391 | -33.8306
23.5747 | -35.4395 | -33.8366
23.5808 | -35.4399 | -33.8425
23.5869 | -35.4403 | -33.8486
23.593 | -35.4407 | -33.8547
23.5991 | -35.4412 | -33.8608
23.6052 | -35.4416 | -33.8669
23.6113 | -35.442 | -33.873
23.6174 | -35.4424 | -33.879
23.6235 | -35.4428 | -33.885
23.6296 | -35.4432 | -33.8911
23.6357 | -35.4436 | -33.897
23.6418 | -35.4441 | -33.9031
23.6479 | -35.4445 | -33.909
23.654 | -35.4449 | -33.9151
23.6601 | -35.4453 | -33.9212
23.6662 | -35.4457 | -33.9272
23.6723 | -35.4461 | -33.9331
23.6784 | -35.4465 | -33.9392
23.6845 | -35.4469 | -33.9451
23.6906 | -35.4474 | -33.9511
23.6967 | -35.4478 | -33.9571
23.7028 | -35.4482 | -33.9631
23.7089 | -35.4486 | -33.9691
23.7151 | -35.449 | -33.9749
23.7212 | -35.4494 | -33.9809
23.7273 | -35.4498 | -33.9869
23.7334 | -35.4503 | -33.993
23.7395 | -35.4507 | -33.9991
23.7456 | -35.4511 | -34.0049
23.7517 | -35.4515 | -34.0109
23.7578 | -35.4519 | -34.0167
23.7639 | -35.4523 | -34.0227
23.77 | -35.4527 | -34.0287
23.7761 | -35.4531 | -34.0348
23.7822 | -35.4535 | -34.0405
23.7883 | -35.4539 | -34.0465
23.7944 | -35.4544 | -34.0525
23.8005 | -35.4548 | -34.0586
23.8066 | -35.4552 | -34.0647
23.8127 | -35.4556 | -34.0707
23.8188 | -35.456 | -34.0767
23.8249 | -35.4564 | -34.0828
23.831 | -35.4569 | -34.0889
23.8371 | -35.4572 | -34.0948
23.8432 | -35.4577 | -34.1008
23.8493 | -35.4581 | -34.1067
23.8555 | -35.4585 | -34.1128
23.8616 | -35.4589 | -34.1187
23.8677 | -35.4593 | -34.1248
23.8738 | -35.4597 | -34.1307
23.8799 | -35.4601 | -34.1368
23.886 | -35.4605 | -34.1428
23.8921 | -35.4609 | -34.1488
23.8982 | -35.4614 | -34.1549
23.9043 | -35.4618 | -34.1611
23.9104 | -35.4622 | -34.1671
23.9165 | -35.4626 | -34.1732
23.9226 | -35.463 | -34.1791
23.9287 | -35.4634 | -34.1852
23.9348 | -35.4638 | -34.1914
23.9409 | -35.4642 | -34.1973
23.947 | -35.4646 | -34.2034
23.9531 | -35.465 | -34.2092
23.9592 | -35.4654 | -34.2152
23.9653 | -35.4658 | -34.2213
23.9714 | -35.4662 | -34.2272
23.9775 | -35.4666 | -34.2332
23.9836 | -35.4671 | -34.2393
23.9897 | -35.4675 | -34.2452
23.9958 | -35.4679 | -34.2513
24.002 | -35.4683 | -34.2573
24.0081 | -35.4687 | -34.2633
24.0142 | -35.4691 | -34.2692
24.0203 | -35.4695 | -34.2752
24.0264 | -35.4699 | -34.2811
24.0325 | -35.4703 | -34.2871
24.0386 | -35.4707 | -34.2931
24.0447 | -35.4711 | -34.2993
24.0508 | -35.4715 | -34.3052
24.0569 | -35.4719 | -34.3112
24.063 | -35.4724 | -34.3171
24.0691 | -35.4728 | -34.323
24.0752 | -35.4732 | -34.3292
24.0813 | -35.4736 | -34.3351
24.0874 | -35.474 | -34.3411
24.0935 | -35.4744 | -34.3471
24.0996 | -35.4748 | -34.3529
24.1057 | -35.4751 | -34.3589
24.1118 | -35.4755 | -34.365
24.1179 | -35.476 | -34.3711
24.124 | -35.4764 | -34.377
24.1301 | -35.4768 | -34.3829
24.1362 | -35.4772 | -34.3889
24.1424 | -35.4776 | -34.395
24.1485 | -35.478 | -34.401
24.1546 | -35.4784 | -34.4071
24.1607 | -35.4788 | -34.4131
24.1668 | -35.4792 | -34.4189
24.1729 | -35.4796 | -34.4249
24.179 | -35.48 | -34.4308
24.1851 | -35.4804 | -34.4367
24.1912 | -35.4808 | -34.4427
24.1973 | -35.4812 | -34.4488
24.2034 | -35.4816 | -34.4547
24.2095 | -35.482 | -34.4609
24.2156 | -35.4824 | -34.4669
24.2217 | -35.4828 | -34.4729
24.2278 | -35.4832 | -34.4787
24.2339 | -35.4836 | -34.4845
24.24 | -35.484 | -34.4903
24.2461 | -35.4845 | -34.4963
24.2522 | -35.4848 | -34.5022
24.2583 | -35.4852 | -34.5081
24.2644 | -35.4856 | -34.5142
24.2705 | -35.486 | -34.5199
24.2766 | -35.4864 | -34.5258
24.2827 | -35.4868 | -34.5316
24.2889 | -35.4872 | -34.5376
24.295 | -35.4876 | -34.5436
24.3011 | -35.488 | -34.5494
24.3072 | -35.4884 | -34.5553
24.3133 | -35.4888 | -34.5612
24.3194 | -35.4892 | -34.5673
24.3255 | -35.4896 | -34.5734
24.3316 | -35.49 | -34.5792
24.3377 | -35.4904 | -34.5851
24.3438 | -35.4907 | -34.5908
24.3499 | -35.4911 | -34.5966
24.356 | -35.4915 | -34.6024
24.3621 | -35.4919 | -34.6084
24.3682 | -35.4923 | -34.6143
24.3743 | -35.4927 | -34.6204
24.3804 | -35.4931 | -34.6262
24.3865 | -35.4935 | -34.6322
24.3926 | -35.4939 | -34.6381
24.3987 | -35.4943 | -34.6442
24.4048 | -35.4947 | -34.6499
24.4109 | -35.4951 | -34.6559
24.417 | -35.4955 | -34.6617
24.4231 | -35.4959 | -34.6677
24.4293 | -35.4963 | -34.6736
24.4354 | -35.4967 | -34.6796
24.4415 | -35.4971 | -34.6854
24.4476 | -35.4975 | -34.6912
24.4537 | -35.4979 | -34.6972
24.4598 | -35.4983 | -34.7032
24.4659 | -35.4987 | -34.7092
24.472 | -35.4991 | -34.7152
24.4781 | -35.4995 | -34.721
24.4842 | -35.4999 | -34.7268
24.4903 | -35.5003 | -34.7328
24.4964 | -35.5007 | -34.7388
24.5025 | -35.5011 | -34.7448
24.5086 | -35.5015 | -34.7507
24.5147 | -35.5019 | -34.7565
24.5208 | -35.5023 | -34.7623
24.5269 | -35.5027 | -34.7682
24.533 | -35.5031 | -34.774
24.5391 | -35.5035 | -34.78
24.5452 | -35.5039 | -34.786
24.5513 | -35.5043 | -34.7917
24.5574 | -35.5047 | -34.7975
24.5635 | -35.5051 | -34.8033
24.5696 | -35.5055 | -34.8093
24.5758 | -35.5059 | -34.8153
24.5819 | -35.5063 | -34.8212
24.588 | -35.5067 | -34.8267
24.5941 | -35.5071 | -34.8326
24.6002 | -35.5075 | -34.8385
24.6063 | -35.5079 | -34.8445
24.6124 | -35.5084 | -34.8504
24.6185 | -35.5088 | -34.8563
24.6246 | -35.5091 | -34.8621
24.6307 | -35.5095 | -34.8681
24.6368 | -35.5099 | -34.874
24.6429 | -35.5103 | -34.8798
24.649 | -35.5107 | -34.8857
24.6551 | -35.5111 | -34.8915
24.6612 | -35.5115 | -34.8971
24.6673 | -35.5119 | -34.903
24.6734 | -35.5123 | -34.909
24.6795 | -35.5127 | -34.915
24.6856 | -35.513 | -34.9209
24.6917 | -35.5135 | -34.9268
24.6978 | -35.5138 | -34.9326
24.7039 | -35.5142 | -34.9385
24.71 | -35.5146 | -34.9445
24.7162 | -35.515 | -34.9505
24.7223 | -35.5154 | -34.9563
24.7284 | -35.5158 | -34.962
24.7345 | -35.5162 | -34.9679
24.7406 | -35.5166 | -34.9738
24.7467 | -35.5169 | -34.9797
24.7528 | -35.5174 | -34.9856
24.7589 | -35.5177 | -34.9914
24.765 | -35.5181 | -34.997
24.7711 | -35.5185 | -35.0028
24.7772 | -35.5189 | -35.0088
24.7833 | -35.5193 | -35.0144
24.7894 | -35.5196 | -35.0203
24.7955 | -35.52 | -35.0261
24.8016 | -35.5204 | -35.0317
24.8077 | -35.5208 | -35.0377
24.8138 | -35.5212 | -35.0437
24.8199 | -35.5216 | -35.0495
24.826 | -35.522 | -35.0554
24.8321 | -35.5224 | -35.0612
24.8382 | -35.5228 | -35.0668
24.8443 | -35.5232 | -35.0727
24.8504 | -35.5235 | -35.0786
24.8565 | -35.5239 | -35.0846
24.8627 | -35.5243 | -35.0904
24.8688 | -35.5247 | -35.0962
24.8749 | -35.5251 | -35.1019
24.881 | -35.5255 | -35.1077
24.8871 | -35.5259 | -35.1137
24.8932 | -35.5263 | -35.1195
24.8993 | -35.5267 | -35.1254
24.9054 | -35.5271 | -35.1311
24.9115 | -35.5275 | -35.137
24.9176 | -35.5279 | -35.143
24.9237 | -35.5283 | -35.1489
24.9298 | -35.5286 | -35.1548
24.9359 | -35.529 | -35.1607
24.942 | -35.5294 | -35.1663
24.9481 | -35.5298 | -35.1722
24.9542 | -35.5302 | -35.1779
24.9603 | -35.5306 | -35.1838
24.9664 | -35.531 | -35.1895
24.9725 | -35.5313 | -35.1955
24.9786 | -35.5317 | -35.201
24.9847 | -35.5321 | -35.2068
24.9908 | -35.5325 | -35.2126
24.9969 | -35.5329 | -35.2183
25.0031 | -35.5333 | -35.2243
25.0092 | -35.5337 | -35.2301
25.0153 | -35.5341 | -35.236
25.0214 | -35.5345 | -35.2417
25.0275 | -35.5348 | -35.2474
25.0336 | -35.5352 | -35.2533
25.0397 | -35.5356 | -35.2593
25.0458 | -35.536 | -35.265
25.0519 | -35.5364 | -35.2708
25.058 | -35.5367 | -35.2765
25.0641 | -35.5371 | -35.2823
25.0702 | -35.5375 | -35.2882
25.0763 | -35.5379 | -35.294
25.0824 | -35.5382 | -35.2995
25.0885 | -35.5386 | -35.3053
25.0946 | -35.539 | -35.3111
25.1007 | -35.5394 | -35.3168
25.1068 | -35.5397 | -35.3225
25.1129 | -35.5401 | -35.3284
25.119 | -35.5405 | -35.334
25.1251 | -35.5409 | -35.3396
25.1312 | -35.5413 | -35.3454
25.1373 | -35.5417 | -35.3512
25.1435 | -35.542 | -35.3572
25.1496 | -35.5424 | -35.3626
25.1557 | -35.5428 | -35.3683
25.1618 | -35.5432 | -35.3741
25.1679 | -35.5436 | -35.38
25.174 | -35.544 | -35.3859
25.1801 | -35.5443 | -35.3917
25.1862 | -35.5448 | -35.3972
25.1923 | -35.5451 | -35.403
25.1984 | -35.5455 | -35.4088
25.2045 | -35.5459 | -35.4147
25.2106 | -35.5463 | -35.4206
25.2167 | -35.5467 | -35.4265
25.2228 | -35.547 | -35.432
25.2289 | -35.5474 | -35.4376
25.235 | -35.5478 | -35.4435
25.2411 | -35.5482 | -35.4492
25.2472 | -35.5485 | -35.455
25.2533 | -35.5489 | -35.4609
25.2594 | -35.5493 | -35.4662
25.2655 | -35.5496 | -35.4721
25.2716 | -35.55 | -35.4779
25.2777 | -35.5504 | -35.4835
25.2838 | -35.5508 | -35.4893
25.29 | -35.5512 | -35.495
25.2961 | -35.5516 | -35.5005
25.3022 | -35.5519 | -35.5063
25.3083 | -35.5523 | -35.5121
25.3144 | -35.5527 | -35.5178
25.3205 | -35.5531 | -35.5237
25.3266 | -35.5534 | -35.5291
25.3327 | -35.5538 | -35.5349
25.3388 | -35.5541 | -35.541
25.3449 | -35.5545 | -35.5467
25.351 | -35.5549 | -35.5525
25.3571 | -35.5553 | -35.5583
25.3632 | -35.5557 | -35.5639
25.3693 | -35.5561 | -35.5697
25.3754 | -35.5565 | -35.5755
25.3815 | -35.5568 | -35.5815
25.3876 | -35.5572 | -35.5872
25.3937 | -35.5576 | -35.593
25.3998 | -35.5579 | -35.5984
25.4059 | -35.5583 | -35.6042
25.412 | -35.5587 | -35.6099
25.4181 | -35.559 | -35.6157
25.4242 | -35.5594 | -35.6215
25.4304 | -35.5598 | -35.6272
25.4365 | -35.5601 | -35.6328
25.4426 | -35.5605 | -35.6386
25.4487 | -35.5609 | -35.6443
25.4548 | -35.5612 | -35.65
25.4609 | -35.5616 | -35.6557
25.467 | -35.562 | -35.6612
25.4731 | -35.5624 | -35.667
25.4792 | -35.5627 | -35.6725
25.4853 | -35.5631 | -35.6784
25.4914 | -35.5635 | -35.684
25.4975 | -35.5639 | -35.6896
25.5036 | -35.5642 | -35.695
25.5097 | -35.5646 | -35.7006
25.5158 | -35.565 | -35.7064
25.5219 | -35.5654 | -35.7123
25.528 | -35.5658 | -35.7181
25.5341 | -35.5662 | -35.7238
25.5402 | -35.5665 | -35.7292
25.5463 | -35.5669 | -35.735
25.5524 | -35.5673 | -35.7406
25.5585 | -35.5676 | -35.7462
25.5646 | -35.568 | -35.7517
25.5707 | -35.5684 | -35.7573
25.5769 | -35.5688 | -35.7629
25.583 | -35.5692 | -35.7685
25.5891 | -35.5696 | -35.7742
25.5952 | -35.5699 | -35.78
25.6013 | -35.5703 | -35.7856
25.6074 | -35.5706 | -35.791
25.6135 | -35.571 | -35.7966
25.6196 | -35.5714 | -35.8025
25.6257 | -35.5717 | -35.8081
25.6318 | -35.5721 | -35.8138
25.6379 | -35.5725 | -35.8194
25.644 | -35.5728 | -35.825
25.6501 | -35.5732 | -35.8305
25.6562 | -35.5736 | -35.836
25.6623 | -35.574 | -35.8417
25.6684 | -35.5744 | -35.8472
25.6745 | -35.5748 | -35.8528
25.6806 | -35.5751 | -35.8581
25.6867 | -35.5754 | -35.8632
25.6928 | -35.5758 | -35.8678
25.6989 | -35.5764 | -35.8731
25.705 | -35.5768 | -35.8786
25.7111 | -35.5772 | -35.8841
25.7173 | -35.5776 | -35.89
25.7234 | -35.578 | -35.8957
25.7295 | -35.5783 | -35.9013
25.7356 | -35.5787 | -35.9069
25.7417 | -35.5791 | -35.9124
25.7478 | -35.5795 | -35.9179
25.7539 | -35.5799 | -35.9234
25.76 | -35.5803 | -35.9292
25.7661 | -35.5806 | -35.9347
25.7722 | -35.581 | -35.9403
25.7783 | -35.5814 | -35.946
25.7844 | -35.5818 | -35.9515
25.7905 | -35.5821 | -35.9572
25.7966 | -35.5824 | -35.9627
25.8027 | -35.5828 | -35.9682
25.8088 | -35.5832 | -35.9737
25.8149 | -35.5835 | -35.9794
25.821 | -35.5839 | -35.9851
25.8271 | -35.5843 | -35.9906
25.8332 | -35.5847 | -35.9962
25.8393 | -35.5851 | -36.0017
25.8454 | -35.5854 | -36.0072
25.8515 | -35.5858 | -36.0126
25.8576 | -35.5861 | -36.0177
25.8638 | -35.5865 | -36.0232
25.8699 | -35.5869 | -36.029
25.876 | -35.5872 | -36.0346
25.8821 | -35.5876 | -36.0401
25.8882 | -35.588 | -36.0455
25.8943 | -35.5884 | -36.0513
25.9004 | -35.5887 | -36.0571
25.9065 | -35.5891 | -36.0629
25.9126 | -35.5895 | -36.0686
25.9187 | -35.5899 | -36.0742
25.9248 | -35.5903 | -36.0796
25.9309 | -35.5907 | -36.0851
25.937 | -35.591 | -36.0906
25.9431 | -35.5914 | -36.0963
25.9492 | -35.5918 | -36.1018
25.9553 | -35.5921 | -36.1075
25.9614 | -35.5925 | -36.113
25.9675 | -35.5929 | -36.1188
25.9736 | -35.5932 | -36.1243
25.9797 | -35.5936 | -36.13
25.9858 | -35.5939 | -36.1354
25.9919 | -35.5943 | -36.1412
25.998 | -35.5947 | -36.1463
26.0042 | -35.595 | -36.1518
26.0103 | -35.5954 | -36.1572
26.0164 | -35.5957 | -36.1628
26.0225 | -35.5961 | -36.1684
26.0286 | -35.5965 | -36.1738
26.0347 | -35.5968 | -36.1792
26.0408 | -35.5972 | -36.1849
26.0469 | -35.5975 | -36.1905
26.053 | -35.5979 | -36.1959
26.0591 | -35.5983 | -36.2016
26.0652 | -35.5986 | -36.2072
26.0713 | -35.599 | -36.2127
26.0774 | -35.5994 | -36.2182
26.0835 | -35.5997 | -36.224
26.0896 | -35.6001 | -36.2292
26.0957 | -35.6004 | -36.2348
26.1018 | -35.6007 | -36.2403
26.1079 | -35.6011 | -36.2459
26.114 | -35.6015 | -36.2516
26.1201 | -35.6018 | -36.257
26.1262 | -35.6022 | -36.2623
26.1323 | -35.6026 | -36.2679
26.1384 | -35.603 | -36.2734
26.1445 | -35.6033 | -36.279
26.1507 | -35.6037 | -36.2847
26.1568 | -35.6041 | -36.2904
26.1629 | -35.6044 | -36.2961
26.169 | -35.6047 | -36.3017
26.1751 | -35.6051 | -36.3069
26.1812 | -35.6055 | -36.3123
26.1873 | -35.6059 | -36.318
26.1934 | -35.6062 | -36.3236
26.1995 | -35.6066 | -36.329
26.2056 | -35.607 | -36.3342
26.2117 | -35.6074 | -36.3398
26.2178 | -35.6077 | -36.3451
26.2239 | -35.6081 | -36.3506
26.23 | -35.6085 | -36.356
26.2361 | -35.6089 | -36.3615
26.2422 | -35.6092 | -36.3667
26.2483 | -35.6096 | -36.3721
26.2544 | -35.61 | -36.3777
26.2605 | -35.6104 | -36.3834
26.2666 | -35.6107 | -36.3889
26.2727 | -35.611 | -36.3943
26.2788 | -35.6114 | -36.3999
26.2849 | -35.6118 | -36.4052
26.2911 | -35.6122 | -36.4105
26.2972 | -35.6125 | -36.416
26.3033 | -35.6129 | -36.4216
26.3094 | -35.6133 | -36.4269
26.3155 | -35.6137 | -36.4323
26.3216 | -35.6141 | -36.4381
26.3277 | -35.6144 | -36.4435
26.3338 | -35.6148 | -36.4487
26.3399 | -35.6151 | -36.4541
26.346 | -35.6155 | -36.4595
26.3521 | -35.6158 | -36.4651
26.3582 | -35.6162 | -36.4705
26.3643 | -35.6165 | -36.4762
26.3704 | -35.6169 | -36.4815
26.3765 | -35.6172 | -36.487
26.3826 | -35.6176 | -36.4925
26.3887 | -35.6179 | -36.498
26.3948 | -35.6183 | -36.5035
26.4009 | -35.6187 | -36.5091
26.407 | -35.619 | -36.5147
26.4131 | -35.6194 | -36.5203
26.4192 | -35.6198 | -36.5255
26.4253 | -35.6201 | -36.5312
26.4314 | -35.6205 | -36.5365
26.4376 | -35.6209 | -36.542
26.4437 | -35.6212 | -36.5474
26.4498 | -35.6216 | -36.5529
26.4559 | -35.622 | -36.5578
26.462 | -35.6223 | -36.5635
26.4681 | -35.6227 | -36.5689
26.4742 | -35.623 | -36.5745
26.4803 | -35.6234 | -36.5797
26.4864 | -35.6237 | -36.585
26.4925 | -35.6241 | -36.5905
26.4986 | -35.6245 | -36.596
26.5047 | -35.6248 | -36.6014
26.5108 | -35.6252 | -36.607
26.5169 | -35.6255 | -36.6124
26.523 | -35.6259 | -36.6174
26.5291 | -35.6262 | -36.6226
26.5352 | -35.6266 | -36.628
26.5413 | -35.6269 | -36.6333
26.5474 | -35.6273 | -36.6388
26.5535 | -35.6276 | -36.6442
26.5596 | -35.628 | -36.6495
26.5657 | -35.6284 | -36.6548
26.5718 | -35.6287 | -36.6604
26.578 | -35.6291 | -36.6658
26.5841 | -35.6294 | -36.6711
26.5902 | -35.6298 | -36.6764
26.5963 | -35.6301 | -36.6815
26.6024 | -35.6305 | -36.6869
26.6085 | -35.6308 | -36.6924
26.6146 | -35.6312 | -36.6978
26.6207 | -35.6315 | -36.703
26.6268 | -35.6319 | -36.7085
26.6329 | -35.6322 | -36.7134
26.639 | -35.6326 | -36.7186
26.6451 | -35.6329 | -36.7242
26.6512 | -35.6333 | -36.7293
26.6573 | -35.6336 | -36.7348
26.6634 | -35.6339 | -36.7398
26.6695 | -35.6343 | -36.7451
26.6756 | -35.6346 | -36.7504
26.6817 | -35.635 | -36.756
26.6878 | -35.6353 | -36.7615
26.6939 | -35.6357 | -36.7668
26.7 | -35.636 | -36.7716
26.7061 | -35.6363 | -36.7768
26.7122 | -35.6367 | -36.7822
26.7183 | -35.6371 | -36.7877
26.7245 | -35.6374 | -36.7932
26.7306 | -35.6377 | -36.7984
26.7367 | -35.6381 | -36.8036
26.7428 | -35.6385 | -36.8088
26.7489 | -35.6388 | -36.8144
26.755 | -35.6391 | -36.8197
26.7611 | -35.6395 | -36.825
26.7672 | -35.6399 | -36.8303
26.7733 | -35.6402 | -36.8352
26.7794 | -35.6405 | -36.8408
26.7855 | -35.6409 | -36.8462
26.7916 | -35.6412 | -36.8513
26.7977 | -35.6415 | -36.8566
26.8038 | -35.6418 | -36.8615
26.8099 | -35.6422 | -36.8668
26.816 | -35.6425 | -36.8723
26.8221 | -35.6429 | -36.8776
26.8282 | -35.6433 | -36.8831
26.8343 | -35.6436 | -36.8886
26.8404 | -35.6439 | -36.8937
26.8465 | -35.6443 | -36.899
26.8526 | -35.6447 | -36.9042
26.8587 | -35.645 | -36.9097
26.8649 | -35.6454 | -36.9153
26.871 | -35.6457 | -36.9204
26.8771 | -35.646 | -36.9254
26.8832 | -35.6464 | -36.9307
26.8893 | -35.6467 | -36.9359
26.8954 | -35.6471 | -36.9412
26.9015 | -35.6474 | -36.9466
26.9076 | -35.6477 | -36.9519
26.9137 | -35.6481 | -36.9568
26.9198 | -35.6484 | -36.9622
26.9259 | -35.6488 | -36.9675
26.932 | -35.6491 | -36.9727
26.9381 | -35.6495 | -36.978
26.9442 | -35.6498 | -36.9836
26.9503 | -35.6501 | -36.9891
26.9564 | -35.6505 | -36.9945
26.9625 | -35.6508 | -36.9999
26.9686 | -35.6511 | -37.0051
26.9747 | -35.6515 | -37.0103
26.9808 | -35.6518 | -37.0154
26.9869 | -35.6521 | -37.0208
26.993 | -35.6524 | -37.026
26.9991 | -35.6528 | -37.0312
27.0052 | -35.6531 | -37.0366
27.0114 | -35.6535 | -37.0419
27.0175 | -35.6538 | -37.047
27.0236 | -35.6541 | -37.0524
27.0297 | -35.6544 | -37.0578
27.0358 | -35.6548 | -37.063
27.0419 | -35.6552 | -37.0683
27.048 | -35.6555 | -37.0735
27.0541 | -35.6559 | -37.0787
27.0602 | -35.6562 | -37.0839
27.0663 | -35.6566 | -37.0893
27.0724 | -35.6569 | -37.0945
27.0785 | -35.6572 | -37.0999
27.0846 | -35.6576 | -37.105
27.0907 | -35.6579 | -37.1102
27.0968 | -35.6582 | -37.1155
27.1029 | -35.6586 | -37.1207
27.109 | -35.659 | -37.1259
27.1151 | -35.6593 | -37.1311
27.1212 | -35.6596 | -37.1362
27.1273 | -35.66 | -37.1413
27.1334 | -35.6603 | -37.1467
27.1395 | -35.6607 | -37.1519
27.1456 | -35.661 | -37.1571
27.1518 | -35.6614 | -37.1624
27.1579 | -35.6617 | -37.1676
27.164 | -35.662 | -37.1728
27.1701 | -35.6624 | -37.178
27.1762 | -35.6627 | -37.1833
27.1823 | -35.663 | -37.1883
27.1884 | -35.6633 | -37.1936
27.1945 | -35.6637 | -37.1986
27.2006 | -35.664 | -37.2038
27.2067 | -35.6644 | -37.209
27.2128 | -35.6647 | -37.2144
27.2189 | -35.665 | -37.2196
27.225 | -35.6654 | -37.225
27.2311 | -35.6657 | -37.2301
27.2372 | -35.6661 | -37.2352
27.2433 | -35.6664 | -37.2404
27.2494 | -35.6668 | -37.2459
27.2555 | -35.6671 | -37.2511
27.2616 | -35.6674 | -37.2561
27.2677 | -35.6678 | -37.2614
27.2738 | -35.6681 | -37.2666
27.2799 | -35.6684 | -37.272
27.286 | -35.6688 | -37.277
27.2921 | -35.6691 | -37.282
27.2983 | -35.6694 | -37.2871
27.3044 | -35.6698 | -37.2924
27.3105 | -35.6701 | -37.2977
27.3166 | -35.6704 | -37.3028
27.3227 | -35.6708 | -37.3079
27.3288 | -35.6711 | -37.3129
27.3349 | -35.6714 | -37.3177
27.341 | -35.6718 | -37.3228
27.3471 | -35.6721 | -37.328
27.3532 | -35.6725 | -37.3332
27.3593 | -35.6728 | -37.3385
27.3654 | -35.6731 | -37.3437
27.3715 | -35.6735 | -37.3485
27.3776 | -35.6738 | -37.354
27.3837 | -35.6742 | -37.3591
27.3898 | -35.6745 | -37.3641
27.3959 | -35.6749 | -37.3693
27.402 | -35.6752 | -37.3741
27.4081 | -35.6755 | -37.3792
27.4142 | -35.6759 | -37.3846
27.4203 | -35.6762 | -37.3899
27.4264 | -35.6765 | -37.3952
27.4325 | -35.6768 | -37.4003
27.4387 | -35.6772 | -37.4054
27.4448 | -35.6775 | -37.4105
27.4509 | -35.6778 | -37.4157
27.457 | -35.6782 | -37.421
27.4631 | -35.6785 | -37.4262
27.4692 | -35.6788 | -37.4314
27.4753 | -35.6791 | -37.4364
27.4814 | -35.6794 | -37.4415
27.4875 | -35.6798 | -37.4467
27.4936 | -35.6801 | -37.4517
27.4997 | -35.6805 | -37.4569
27.5058 | -35.6808 | -37.4622
27.5119 | -35.6811 | -37.4671
27.518 | -35.6815 | -37.4722
27.5241 | -35.6818 | -37.4774
27.5302 | -35.6821 | -37.4827
27.5363 | -35.6824 | -37.4878
27.5424 | -35.6827 | -37.4927
27.5485 | -35.6831 | -37.4978
27.5546 | -35.6834 | -37.5029
27.5607 | -35.6838 | -37.5081
27.5668 | -35.6841 | -37.5133
27.5729 | -35.6844 | -37.5183
27.5791 | -35.6847 | -37.5235
27.5852 | -35.685 | -37.5286
27.5913 | -35.6854 | -37.5337
27.5974 | -35.6857 | -37.5389
27.6035 | -35.686 | -37.5441
27.6096 | -35.6863 | -37.5492
27.6157 | -35.6866 | -37.5541
27.6218 | -35.6869 | -37.5593
27.6279 | -35.6872 | -37.5644
27.634 | -35.6876 | -37.5695
27.6401 | -35.6879 | -37.5746
27.6462 | -35.6882 | -37.5797
27.6523 | -35.6885 | -37.5846
27.6584 | -35.6888 | -37.5896
27.6645 | -35.6892 | -37.5947
27.6706 | -35.6895 | -37.5997
27.6767 | -35.6898 | -37.6049
27.6828 | -35.6901 | -37.6099
27.6889 | -35.6904 | -37.6151
27.695 | -35.6907 | -37.6201
27.7011 | -35.691 | -37.6252
27.7072 | -35.6914 | -37.6302
27.7133 | -35.6917 | -37.6355
27.7194 | -35.692 | -37.6403
27.7256 | -35.6924 | -37.6454
27.7317 | -35.6927 | -37.6505
27.7378 | -35.693 | -37.6556
27.7439 | -35.6933 | -37.6609
27.75 | -35.6937 | -37.666
27.7561 | -35.694 | -37.6708
27.7622 | -35.6943 | -37.6759
27.7683 | -35.6946 | -37.681
27.7744 | -35.695 | -37.6861
27.7805 | -35.6953 | -37.6912
27.7866 | -35.6956 | -37.6964
27.7927 | -35.6959 | -37.7013
27.7988 | -35.6962 | -37.7064
27.8049 | -35.6965 | -37.7115
27.811 | -35.6968 | -37.7164
27.8171 | -35.6972 | -37.7215
27.8232 | -35.6975 | -37.7264
27.8293 | -35.6978 | -37.7314
27.8354 | -35.6982 | -37.7366
27.8415 | -35.6985 | -37.7416
27.8476 | -35.6988 | -37.7468
27.8537 | -35.6991 | -37.7519
27.8598 | -35.6994 | -37.7569
27.866 | -35.6998 | -37.7618
27.8721 | -35.7001 | -37.7668
27.8782 | -35.7004 | -37.7719
27.8843 | -35.7007 | -37.7769
27.8904 | -35.7011 | -37.7821
27.8965 | -35.7014 | -37.7869
27.9026 | -35.7017 | -37.792
27.9087 | -35.702 | -37.7972
27.9148 | -35.7023 | -37.8023
27.9209 | -35.7026 | -37.8074
27.927 | -35.703 | -37.8124
27.9331 | -35.7033 | -37.8172
27.9392 | -35.7036 | -37.8224
27.9453 | -35.7039 | -37.8274
27.9514 | -35.7042 | -37.8325
27.9575 | -35.7045 | -37.8375
27.9636 | -35.7048 | -37.8425
27.9697 | -35.7051 | -37.8475
27.9758 | -35.7054 | -37.8526
27.9819 | -35.7057 | -37.8579
27.988 | -35.706 | -37.8629
27.9941 | -35.7064 | -37.868
28.0002 | -35.7067 | -37.8728
28.0063 | -35.707 | -37.878
28.0125 | -35.7073 | -37.8831
28.0186 | -35.7076 | -37.8883
28.0247 | -35.708 | -37.8935
28.0308 | -35.7083 | -37.8985
28.0369 | -35.7086 | -37.9033
28.043 | -35.7089 | -37.9084
28.0491 | -35.7092 | -37.9133
28.0552 | -35.7095 | -37.9185
28.0613 | -35.7098 | -37.9235
28.0674 | -35.7101 | -37.9286
28.0735 | -35.7104 | -37.9335

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4131 | -42.7388
21.4908 | -35.4952 | -34.8053
16.4948 | -34.891 | -27.4682
12.6602 | -34.5358 | -21.2422
9.7171 | -34.323 | -16.3749
7.4582 | -34.1782 | -12.4216
5.7244 | -34.0725 | -9.2274
4.3936 | -34.0075 | -6.6413
3.3722 | -33.9889 | -4.608
2.5883 | -33.989 | -3.0685
1.9866 | -33.9819 | -1.8372
1.5248 | -33.9859 | -0.85145
1.1703 | -33.9919 | -0.20994
1 | -33.9962 | 0.10951
0.89822 | -33.9834 | 0.37793
0.68941 | -33.9674 | 0.95196
0.52919 | -33.9828 | 1.3957
0.40612 | -33.9719 | 2.0002
0.31171 | -33.9637 | 2.4461
0.23928 | -33.9632 | 3.1665
0.18358 | -33.9875 | 3.5773
0.14091 | -33.9842 | 4.4587
0.10819 | -34.0072 | 5.6076
0.083017 | -34.0396 | 7.0757
0.063705 | -34.0729 | 8.9377
0.048923 | -34.1386 | 11.4151
0.037479 | -34.2628 | 14.5887
0.028801 | -34.4345 | 18.6062
0.022078 | -34.7381 | 23.9928
0.017023 | -35.216 | 30.4386
0.013018 | -36.0216 | 38.13
0.010014 | -37.1378 | 46.0404

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003704 | -9.8887
0.007407 | -26.7119
0.011111 | -34.7962
0.014815 | -39.2429
0.018519 | -40.9551
0.022222 | -43.4994
0.025926 | -45.4671
0.02963 | -46.5393
0.033333 | -48.2602
0.037037 | -48.7838
0.040741 | -50.69
0.044444 | -52.0751
0.048148 | -52.4588
0.051852 | -53.9231
0.055556 | -54.6763
0.059259 | -55.0954
0.062963 | -55.9125
0.066667 | -55.8188
0.07037 | -57.019
0.074074 | -58.2017
0.077778 | -59.5871
0.081481 | -59.6096
0.085185 | -59.6134
0.088889 | -59.3785
0.092593 | -60.2128
0.096296 | -61.0066
0.1 | -62.1558
0.1037 | -63.2106
0.10741 | -62.8772
0.11111 | -63.1509
0.11482 | -63.5225
0.11852 | -63.6362
0.12222 | -63.4543
0.12593 | -64.802
0.12963 | -65.0427
0.13333 | -65.5446
0.13704 | -65.7662
0.14074 | -65.8992
0.14444 | -65.4646
0.14815 | -67.53
0.15185 | -67.7036
0.15556 | -66.9314
0.15926 | -67.3466
0.16296 | -67.9106
0.16667 | -68.1518
0.17037 | -67.6644
0.17407 | -68.1582
0.17778 | -68.9403
0.18148 | -69.0243
0.18518 | -69.2332
0.18889 | -69.4135
0.19259 | -70.1216
0.1963 | -69.7546
0.2 | -69.2148
0.2037 | -69.5728
0.20741 | -70.3588
0.21111 | -71.1125
0.21482 | -71.1321
0.21852 | -71.1172
0.22222 | -71.3596
0.22593 | -71.2673
0.22963 | -71.3758
0.23333 | -71.8478
0.23704 | -71.8462
0.24074 | -71.9734
0.24444 | -72.8907
0.24815 | -72.9179
0.25185 | -73.0704
0.25556 | -74.3119
0.25926 | -74.1894
0.26296 | -73.3441
0.26667 | -73.2225
0.27037 | -73.8599
0.27407 | -73.3616
0.27778 | -73.3378
0.28148 | -74.0028
0.28519 | -73.8654
0.28889 | -74.11
0.29259 | -74.8785
0.2963 | -75.6883
0.3 | -75.9066
0.3037 | -75.7207
0.30741 | -75.8538
0.31111 | -76.0039
0.31482 | -75.4368
0.31852 | -75.0312
0.32222 | -75.518
0.32593 | -75.9111
0.32963 | -76.2876
0.33333 | -77.5596
0.33704 | -76.8257
0.34074 | -77.1413
0.34444 | -76.7954
0.34815 | -76.5648
0.35185 | -75.9068
0.35556 | -75.5931
0.35926 | -76.4732
0.36296 | -77.7313
0.36667 | -76.6854
0.37037 | -77.0417
0.37407 | -77.2703
0.37778 | -77.243
0.38148 | -77.0361
0.38518 | -77.1758
0.38889 | -76.792
0.39259 | -77.2747
0.3963 | -77.6567
0.4 | -78.2792
0.4037 | -77.8961
0.40741 | -77.656
0.41111 | -76.8983
0.41481 | -77.4213
0.41852 | -78.6253
0.42222 | -77.9617
0.42593 | -78.0811
0.42963 | -78.8997
0.43333 | -79.0187
0.43704 | -78.5953
0.44074 | -78.532
0.44444 | -78.0525
0.44815 | -77.148
0.45185 | -77.4744
0.45556 | -79.218
0.45926 | -80.4651
0.46296 | -80.7407
0.46667 | -80.4737
0.47037 | -80.3075
0.47407 | -79.8922
0.47778 | -79.6978
0.48148 | -79.6527
0.48518 | -79.7789
0.48889 | -80.1521
0.49259 | -80.528
0.4963 | -79.9952
0.5 | -80.456
0.5037 | -79.732
0.50741 | -80.1053
0.51111 | -80.381
0.51482 | -80.0903
0.51852 | -80.0764
0.52222 | -80.2966
0.52593 | -79.7511
0.52963 | -79.7934
0.53333 | -79.7254
0.53704 | -80.0946
0.54074 | -80.3272
0.54444 | -80.3606
0.54815 | -80.4391
0.55185 | -80.2386
0.55556 | -80.3343
0.55926 | -81.1161
0.56296 | -81.9106
0.56667 | -82.4168
0.57037 | -82.0504
0.57407 | -81.4582
0.57778 | -81.7355
0.58148 | -82.5679
0.58518 | -82.8526
0.58889 | -82.4024
0.59259 | -82.0314
0.5963 | -82.3287
0.6 | -82.5745
0.6037 | -82.586
0.60741 | -82.1808
0.61111 | -81.7318
0.61482 | -81.5444
0.61852 | -81.4856
0.62222 | -81.6203
0.62593 | -81.7094
0.62963 | -81.8974
0.63333 | -82.4561
0.63704 | -82.465
0.64074 | -82.0078
0.64444 | -82.0645
0.64815 | -82.4002
0.65185 | -82.6859
0.65556 | -82.7327
0.65926 | -82.5972
0.66296 | -82.392
0.66667 | -82.1904
0.67037 | -82.26
0.67407 | -82.9541
0.67778 | -83.4663
0.68148 | -83.2396
0.68519 | -82.7283
0.68889 | -82.398
0.69259 | -82.4537
0.6963 | -82.8042
0.7 | -83.6332
0.7037 | -84.3344
0.70741 | -84.1053
0.71111 | -83.8044
0.71481 | -83.9561
0.71852 | -83.8187
0.72222 | -83.3741
0.72593 | -83.3556
0.72963 | -83.7367
0.73333 | -84.1075
0.73704 | -83.9861
0.74074 | -83.7625
0.74444 | -83.9014
0.74815 | -84.0124
0.75185 | -84.0496
0.75556 | -84.3081
0.75926 | -84.7956
0.76296 | -84.848
0.76667 | -84.754
0.77037 | -84.655
0.77407 | -84.5698
0.77778 | -84.8164
0.78148 | -85.0182
0.78519 | -84.3238
0.78889 | -83.7736
0.79259 | -83.9602
0.7963 | -84.1701
0.8 | -84.6437
0.8037 | -85.0885
0.80741 | -85.0312
0.81111 | -84.958
0.81481 | -85.1499
0.81852 | -85.1855
0.82222 | -85.2757
0.82593 | -85.1494
0.82963 | -84.9948
0.83333 | -85.3841
0.83704 | -85.3939
0.84074 | -85.0801
0.84444 | -85.5304
0.84815 | -85.898
0.85185 | -85.428
0.85556 | -85.295
0.85926 | -85.5578
0.86296 | -85.2037
0.86667 | -84.6477
0.87037 | -84.4049
0.87407 | -84.9398
0.87778 | -85.4865
0.88148 | -85.6383
0.88518 | -85.8814
0.88889 | -86.4114
0.89259 | -86.9852
0.8963 | -87.5293
0.9 | -87.5172
0.9037 | -86.7306
0.90741 | -85.8524
0.91111 | -85.5562
0.91482 | -86.0357
0.91852 | -86.9609
0.92222 | -87.8744
0.92593 | -88.1932
0.92963 | -87.4583
0.93333 | -86.371
0.93704 | -85.7466
0.94074 | -85.4876
0.94444 | -85.7199
0.94815 | -86.4077
0.95185 | -86.8142
0.95556 | -87.0054
0.95926 | -86.9734
0.96296 | -86.9656
0.96667 | -86.8785
0.97037 | -86.8836
0.97407 | -87.321
0.97778 | -87.5731
0.98148 | -87.3595
0.98518 | -87.1173
0.98889 | -86.9063
0.99259 | -86.553
0.9963 | -86.6893
1 | -86.9831
1.0037 | -86.9055
1.0074 | -86.9051
1.0111 | -87.1668
1.0148 | -87.164
1.0185 | -86.943
1.0222 | -86.6517
1.0259 | -86.6983
1.0296 | -87.0865
1.0333 | -87.2465
1.037 | -86.6882
1.0407 | -86.3964
1.0444 | -87.0709
1.0481 | -87.8985
1.0519 | -88.4326
1.0556 | -88.8145
1.0593 | -88.8375
1.063 | -88.5581
1.0667 | -88.2231
1.0704 | -88.1953
1.0741 | -88.4231
1.0778 | -88.5916
1.0815 | -88.4265
1.0852 | -88.0449
1.0889 | -87.752
1.0926 | -87.6033
1.0963 | -87.4865
1.1 | -87.4686
1.1037 | -87.6294
1.1074 | -87.9608
1.1111 | -88.3219
1.1148 | -88.4079
1.1185 | -88.2913
1.1222 | -88.3966
1.1259 | -88.781
1.1296 | -89.1965
1.1333 | -89.4724
1.137 | -89.4074
1.1407 | -88.8421
1.1444 | -88.2184
1.1481 | -88.0351
1.1519 | -88.2624
1.1556 | -88.5662
1.1593 | -88.6296
1.163 | -88.4599
1.1667 | -88.4176
1.1704 | -88.6134
1.1741 | -88.7529
1.1778 | -88.6279
1.1815 | -88.3951
1.1852 | -88.2947
1.1889 | -88.37
1.1926 | -88.489
1.1963 | -88.6456
1.2 | -88.9198
1.2037 | -89.1795
1.2074 | -89.1643
1.2111 | -88.8301
1.2148 | -88.4444
1.2185 | -88.3362
1.2222 | -88.4909
1.2259 | -88.6228
1.2296 | -88.7063
1.2333 | -88.947
1.237 | -89.32
1.2407 | -89.6855
1.2444 | -89.916
1.2481 | -89.9147
1.2519 | -89.8366
1.2556 | -89.8825
1.2593 | -89.9969
1.263 | -89.9672
1.2667 | -89.7952
1.2704 | -89.754
1.2741 | -89.9205
1.2778 | -90.2031
1.2815 | -90.5461
1.2852 | -90.7805
1.2889 | -90.7135
1.2926 | -90.4852
1.2963 | -90.3419
1.3 | -90.1777
1.3037 | -89.8258
1.3074 | -89.5361
1.3111 | -89.491
1.3148 | -89.5294
1.3185 | -89.5848
1.3222 | -89.7385
1.3259 | -90.0289
1.3296 | -90.3862
1.3333 | -90.498
1.337 | -90.1869
1.3407 | -89.8023
1.3444 | -89.6297
1.3481 | -89.6975
1.3519 | -89.8869
1.3556 | -89.9639
1.3593 | -89.9234
1.363 | -89.9126
1.3667 | -89.7106
1.3704 | -89.336
1.3741 | -89.2778
1.3778 | -89.7019
1.3815 | -90.3534
1.3852 | -90.8682
1.3889 | -90.9423
1.3926 | -90.6105
1.3963 | -90.2634
1.4 | -90.1819
1.4037 | -90.279
1.4074 | -90.2657
1.4111 | -90.1335
1.4148 | -90.0238
1.4185 | -89.9777
1.4222 | -90.0291
1.4259 | -90.1267
1.4296 | -90.0779
1.4333 | -89.8869
1.437 | -89.8997
1.4407 | -90.2829
1.4444 | -90.7069
1.4481 | -90.8172
1.4519 | -90.6605
1.4556 | -90.4875
1.4593 | -90.4443
1.463 | -90.4234
1.4667 | -90.3766
1.4704 | -90.459
1.4741 | -90.6894
1.4778 | -90.8238
1.4815 | -90.6998
1.4852 | -90.4933
1.4889 | -90.43
1.4926 | -90.5325
1.4963 | -90.8131
1.5 | -91.2835
1.5037 | -91.7722
1.5074 | -92.0634
1.5111 | -92.1499
1.5148 | -92.1744
1.5185 | -92.1445
1.5222 | -91.9985
1.5259 | -91.8621
1.5296 | -91.8511
1.5333 | -91.88
1.537 | -91.7377
1.5407 | -91.4427
1.5444 | -91.2883
1.5481 | -91.3841
1.5519 | -91.5068
1.5556 | -91.4849
1.5593 | -91.5035
1.563 | -91.6183
1.5667 | -91.5155
1.5704 | -91.2449
1.5741 | -91.1581
1.5778 | -91.3002
1.5815 | -91.5508
1.5852 | -91.8199
1.5889 | -91.9122
1.5926 | -91.7179
1.5963 | -91.3749
1.6 | -91.2281
1.6037 | -91.2021
1.6074 | -91.3207
1.6111 | -91.4712
1.6148 | -91.5347
1.6185 | -91.5316
1.6222 | -91.6001
1.6259 | -91.7809
1.6296 | -91.964
1.6333 | -91.9941
1.637 | -91.8672
1.6407 | -91.6705
1.6444 | -91.5041
1.6481 | -91.4244
1.6519 | -91.4393
1.6556 | -91.4854
1.6593 | -91.5103
1.663 | -91.5487
1.6667 | -91.6591
1.6704 | -91.8009
1.6741 | -91.8385
1.6778 | -91.7135
1.6815 | -91.5353
1.6852 | -91.4409
1.6889 | -91.5168
1.6926 | -91.7734
1.6963 | -92.0808
1.7 | -92.235
1.7037 | -92.1852
1.7074 | -92.1171
1.7111 | -92.1512
1.7148 | -92.2296
1.7185 | -92.2666
1.7222 | -92.2778
1.7259 | -92.3514
1.7296 | -92.5084
1.7333 | -92.6496
1.737 | -92.6253
1.7407 | -92.4232
1.7444 | -92.1754
1.7481 | -91.9677
1.7519 | -91.8565
1.7556 | -91.9
1.7593 | -92.1132
1.763 | -92.3901
1.7667 | -92.5229
1.7704 | -92.4596
1.7741 | -92.3761
1.7778 | -92.4232
1.7815 | -92.5645
1.7852 | -92.6606
1.7889 | -92.6531
1.7926 | -92.5836
1.7963 | -92.5452
1.8 | -92.6102
1.8037 | -92.7122
1.8074 | -92.71
1.8111 | -92.569
1.8148 | -92.4034
1.8185 | -92.3018
1.8222 | -92.274
1.8259 | -92.2761
1.8296 | -92.2693
1.8333 | -92.27
1.837 | -92.3029
1.8407 | -92.3332
1.8444 | -92.3017
1.8481 | -92.1994
1.8519 | -92.1049
1.8556 | -92.0937
1.8593 | -92.1335
1.863 | -92.1185
1.8667 | -92.0283
1.8704 | -91.9824
1.8741 | -92.0695
1.8778 | -92.2624
1.8815 | -92.4791
1.8852 | -92.6766
1.8889 | -92.8327
1.8926 | -92.9708
1.8963 | -93.1139
1.9 | -93.2554
1.9037 | -93.332
1.9074 | -93.2568
1.9111 | -93.0899
1.9148 | -92.9968
1.9185 | -93.0779
1.9222 | -93.2773
1.9259 | -93.4188
1.9296 | -93.3617
1.9333 | -93.0678
1.937 | -92.6531
1.9407 | -92.3351
1.9444 | -92.244
1.9481 | -92.3885
1.9519 | -92.6586
1.9556 | -92.9282
1.9593 | -93.1145
1.963 | -93.1872
1.9667 | -93.1888
1.9704 | -93.1998
1.9741 | -93.2723
1.9778 | -93.3534
1.9815 | -93.3467
1.9852 | -93.2628
1.9889 | -93.1752
1.9926 | -93.1253
1.9963 | -93.1463
2 | -93.2722
2.0037 | -93.4728
2.0074 | -93.5939
2.0111 | -93.493
2.0148 | -93.2368
2.0185 | -93.0355
2.0222 | -93.0205
2.0259 | -93.1664
2.0296 | -93.3482
2.0333 | -93.473
2.037 | -93.5231
2.0407 | -93.4814
2.0444 | -93.3596
2.0481 | -93.2157
2.0519 | -93.1302
2.0556 | -93.1495
2.0593 | -93.2716
2.063 | -93.4473
2.0667 | -93.5926
2.0704 | -93.6715
2.0741 | -93.7484
2.0778 | -93.8468
2.0815 | -93.9022
2.0852 | -93.8434
2.0889 | -93.716
2.0926 | -93.6319
2.0963 | -93.6425
2.1 | -93.7117
2.1037 | -93.807
2.1074 | -93.9574
2.1111 | -94.1777
2.1148 | -94.3665
2.1185 | -94.3614
2.1222 | -94.163
2.1259 | -93.9427
2.1296 | -93.8618
2.1333 | -93.9526
2.137 | -94.1952
2.1407 | -94.4072
2.1444 | -94.6197
2.1481 | -94.7724
2.1519 | -94.8254
2.1556 | -94.7855
2.1593 | -94.6769
2.163 | -94.5332
2.1667 | -94.3845
2.1704 | -94.2671
2.1741 | -94.2037
2.1778 | -94.1895
2.1815 | -94.2183
2.1852 | -94.2527
2.1889 | -94.2332
2.1926 | -94.133
2.1963 | -93.996
2.2 | -93.904
2.2037 | -93.9198
2.2074 | -94.024
2.2111 | -94.1459
2.2148 | -94.2111
2.2185 | -94.1797
2.2222 | -94.0795
2.2259 | -93.9832
2.2296 | -93.9448
2.2333 | -93.9914
2.237 | -94.1107
2.2407 | -94.2781
2.2444 | -94.4241
2.2481 | -94.4916
2.2519 | -94.4682
2.2556 | -94.4024
2.2593 | -94.3676
2.263 | -94.3976
2.2667 | -94.4955
2.2704 | -94.6212
2.2741 | -94.7173
2.2778 | -94.7593
2.2815 | -94.7271
2.2852 | -94.6561
2.2889 | -94.5858
2.2926 | -94.5651
2.2963 | -94.592
2.3 | -94.6148
2.3037 | -94.5795
2.3074 | -94.4478
2.3111 | -94.2693
2.3148 | -94.1141
2.3185 | -94.0098
2.3222 | -93.9957
2.3259 | -94.0762
2.3296 | -94.2325
2.3333 | -94.41
2.337 | -94.5533
2.3407 | -94.6107
2.3444 | -94.5836
2.3481 | -94.5234
2.3519 | -94.4506
2.3556 | -94.3942
2.3593 | -94.3978
2.363 | -94.4791
2.3667 | -94.6491
2.3704 | -94.8666
2.3741 | -95.0545
2.3778 | -95.1677
2.3815 | -95.2238
2.3852 | -95.2267
2.3889 | -95.1825
2.3926 | -95.1043
2.3963 | -95.027
2.4 | -94.991
2.4037 | -95.0229
2.4074 | -95.0959
2.4111 | -95.1379
2.4148 | -95.1064
2.4185 | -95.0205
2.4222 | -94.9394
2.4259 | -94.9134
2.4296 | -94.9342
2.4333 | -94.9655
2.437 | -94.9753
2.4407 | -94.9754
2.4444 | -94.9914
2.4481 | -95.0406
2.4519 | -95.1142
2.4556 | -95.1714
2.4593 | -95.1685
2.463 | -95.0841
2.4667 | -94.9209
2.4704 | -94.734
2.4741 | -94.6087
2.4778 | -94.6125
2.4815 | -94.7689
2.4852 | -95.0383
2.4889 | -95.3387
2.4926 | -95.5244
2.4963 | -95.4662
2.5 | -95.1945
2.5037 | -94.8321
2.5074 | -94.5377
2.5111 | -94.3909
2.5148 | -94.4134
2.5185 | -94.5735
2.5222 | -94.7918
2.5259 | -94.9915
2.5296 | -95.097
2.5333 | -95.0786
2.537 | -94.9573
2.5407 | -94.7878
2.5444 | -94.6898
2.5481 | -94.7095
2.5519 | -94.8172
2.5556 | -94.9446
2.5593 | -95.0297
2.563 | -95.0525
2.5667 | -95.024
2.5704 | -94.9687
2.5741 | -94.9053
2.5778 | -94.852
2.5815 | -94.8353
2.5852 | -94.8605
2.5889 | -94.9301
2.5926 | -95.0162
2.5963 | -95.0813
2.6 | -95.0945
2.6037 | -95.0729
2.6074 | -95.0593
2.6111 | -95.1109
2.6148 | -95.2545
2.6185 | -95.4619
2.6222 | -95.6557
2.6259 | -95.7561
2.6296 | -95.7543
2.6333 | -95.6972
2.637 | -95.6157
2.6407 | -95.5287
2.6444 | -95.4103
2.6481 | -95.2797
2.6519 | -95.1906
2.6556 | -95.1812
2.6593 | -95.238
2.663 | -95.3153
2.6667 | -95.3775
2.6704 | -95.3762
2.6741 | -95.3935
2.6778 | -95.4082
2.6815 | -95.4256
2.6852 | -95.4506
2.6889 | -95.4932
2.6926 | -95.5416
2.6963 | -95.5809
2.7 | -95.6046
2.7037 | -95.6065
2.7074 | -95.5959
2.7111 | -95.5646
2.7148 | -95.5148
2.7185 | -95.4395
2.7222 | -95.3488
2.7259 | -95.2624
2.7296 | -95.1961
2.7333 | -95.1464
2.737 | -95.1254
2.7407 | -95.1478
2.7444 | -95.2292
2.7481 | -95.3689
2.7519 | -95.5386
2.7556 | -95.6857
2.7593 | -95.7734
2.763 | -95.7706
2.7667 | -95.6815
2.7704 | -95.5488
2.7741 | -95.4304
2.7778 | -95.3599
2.7815 | -95.372
2.7852 | -95.448
2.7889 | -95.5598
2.7926 | -95.6558
2.7963 | -95.7113
2.8 | -95.7146
2.8037 | -95.6841
2.8074 | -95.635
2.8111 | -95.6015
2.8148 | -95.6151
2.8185 | -95.6829
2.8222 | -95.8024
2.8259 | -95.9804
2.8296 | -96.1813
2.8333 | -96.3771
2.837 | -96.5187
2.8407 | -96.5742
2.8444 | -96.5166
2.8481 | -96.3707
2.8519 | -96.1809
2.8556 | -96.0076
2.8593 | -95.8842
2.863 | -95.826
2.8667 | -95.8315
2.8704 | -95.8813
2.8741 | -95.9465
2.8778 | -96.0353
2.8815 | -96.1614
2.8852 | -96.337
2.8889 | -96.5215
2.8926 | -96.668
2.8963 | -96.721
2.9 | -96.6688
2.9037 | -96.532
2.9074 | -96.3717
2.9111 | -96.2474
2.9148 | -96.2088
2.9185 | -96.2892
2.9222 | -96.4651
2.9259 | -96.6573
2.9296 | -96.7653
2.9333 | -96.737
2.937 | -96.5892
2.9407 | -96.3777
2.9444 | -96.1716
2.9481 | -96.012
2.9519 | -95.9272
2.9556 | -95.9248
2.9593 | -96.0111
2.963 | -96.1519
2.9667 | -96.3028
2.9704 | -96.4132
2.9741 | -96.4462
2.9778 | -96.3987
2.9815 | -96.3006
2.9852 | -96.2068
2.9889 | -96.1567
2.9926 | -96.1605
2.9963 | -96.2224
3 | -96.3018
3.0037 | -96.3674
3.0074 | -96.3934
3.0111 | -96.3729
3.0148 | -96.3129
3.0185 | -96.2533
3.0222 | -96.2122
3.0259 | -96.1995
3.0296 | -96.2191
3.0333 | -96.2723
3.037 | -96.3595
3.0407 | -96.468
3.0444 | -96.5636
3.0481 | -96.6449
3.0519 | -96.7007
3.0556 | -96.7539
3.0593 | -96.8085
3.063 | -96.832
3.0667 | -96.8071
3.0704 | -96.7386
3.0741 | -96.6453
3.0778 | -96.544
3.0815 | -96.4687
3.0852 | -96.4488
3.0889 | -96.5014
3.0926 | -96.6215
3.0963 | -96.771
3.1 | -96.9187
3.1037 | -97.0332
3.1074 | -97.1102
3.1111 | -97.1354
3.1148 | -97.0912
3.1185 | -96.9718
3.1222 | -96.8069
3.1259 | -96.6426
3.1296 | -96.5133
3.1333 | -96.4481
3.137 | -96.4519
3.1407 | -96.5109
3.1444 | -96.6266
3.1481 | -96.7797
3.1519 | -96.9476
3.1556 | -97.1154
3.1593 | -97.2827
3.163 | -97.4451
3.1667 | -97.6157
3.1704 | -97.7693
3.1741 | -97.8803
3.1778 | -97.9097
3.1815 | -97.8356
3.1852 | -97.6842
3.1889 | -97.518
3.1926 | -97.3646
3.1963 | -97.2635
3.2 | -97.2237
3.2037 | -97.2282
3.2074 | -97.2079
3.2111 | -97.195
3.2148 | -97.1951
3.2185 | -97.2065
3.2222 | -97.235
3.2259 | -97.2573
3.2296 | -97.2679
3.2333 | -97.2543
3.237 | -97.2164
3.2407 | -97.1638
3.2444 | -97.1157
3.2481 | -97.078
3.2519 | -97.0542
3.2556 | -97.03
3.2593 | -97.008
3.263 | -96.9863
3.2667 | -96.9544
3.2704 | -96.9244
3.2741 | -96.9132
3.2778 | -96.9228
3.2815 | -96.9587
3.2852 | -97.003
3.2889 | -97.0624
3.2926 | -97.137
3.2963 | -97.2068
3.3 | -97.2789
3.3037 | -97.3435
3.3074 | -97.3846
3.3111 | -97.4027
3.3148 | -97.3912
3.3185 | -97.3479
3.3222 | -97.2861
3.3259 | -97.2369
3.3296 | -97.1942
3.3333 | -97.1625
3.337 | -97.1428
3.3407 | -97.142
3.3444 | -97.1532
3.3481 | -97.1768
3.3519 | -97.2195
3.3556 | -97.2628
3.3593 | -97.2858
3.363 | -97.2781
3.3667 | -97.2417
3.3704 | -97.1903
3.3741 | -97.148
3.3778 | -97.1299
3.3815 | -97.1365
3.3852 | -97.1912
3.3889 | -97.2883
3.3926 | -97.4036
3.3963 | -97.5225
3.4 | -97.6135
3.4037 | -97.6555
3.4074 | -97.628
3.4111 | -97.5521
3.4148 | -97.4699
3.4185 | -97.4164
3.4222 | -97.4066
3.4259 | -97.4394
3.4296 | -97.4917
3.4333 | -97.5256
3.437 | -97.524
3.4407 | -97.4868
3.4444 | -97.4352
3.4481 | -97.3953
3.4519 | -97.3802
3.4556 | -97.4005
3.4593 | -97.4452
3.463 | -97.4894
3.4667 | -97.5104
3.4704 | -97.5039
3.4741 | -97.4827
3.4778 | -97.4598
3.4815 | -97.4511
3.4852 | -97.4706
3.4889 | -97.524
3.4926 | -97.5941
3.4963 | -97.6708
3.5 | -97.7258
3.5037 | -97.7502
3.5074 | -97.7347
3.5111 | -97.6862
3.5148 | -97.6388
3.5185 | -97.6165
3.5222 | -97.6261
3.5259 | -97.6713
3.5296 | -97.7432
3.5333 | -97.8306
3.537 | -97.9286
3.5407 | -98.0318
3.5444 | -98.1228
3.5481 | -98.1843
3.5519 | -98.2066
3.5556 | -98.1943
3.5593 | -98.152
3.563 | -98.1117
3.5667 | -98.0929
3.5704 | -98.0905
3.5741 | -98.1104
3.5778 | -98.1445
3.5815 | -98.1835
3.5852 | -98.2201
3.5889 | -98.2428
3.5926 | -98.2355
3.5963 | -98.2044
3.6 | -98.1593
3.6037 | -98.1175
3.6074 | -98.0729
3.6111 | -98.0235
3.6148 | -97.9796
3.6185 | -97.9537
3.6222 | -97.9674
3.6259 | -98.0286
3.6296 | -98.1463
3.6333 | -98.3068
3.637 | -98.4717
3.6407 | -98.6069
3.6444 | -98.7022
3.6481 | -98.7325
3.6519 | -98.6926
3.6556 | -98.5997
3.6593 | -98.4728
3.663 | -98.3531
3.6667 | -98.2543
3.6704 | -98.1802
3.6741 | -98.1376
3.6778 | -98.1275
3.6815 | -98.1352
3.6852 | -98.1448
3.6889 | -98.1461
3.6926 | -98.1472
3.6963 | -98.156
3.7 | -98.1727
3.7037 | -98.2081
3.7074 | -98.2507
3.7111 | -98.2859
3.7148 | -98.2993
3.7185 | -98.2909
3.7222 | -98.2769
3.7259 | -98.2719
3.7296 | -98.2801
3.7333 | -98.2995
3.737 | -98.3233
3.7407 | -98.3388
3.7444 | -98.3417
3.7481 | -98.3426
3.7519 | -98.3375
3.7556 | -98.3195
3.7593 | -98.2915
3.763 | -98.251
3.7667 | -98.1907
3.7704 | -98.1152
3.7741 | -98.0307
3.7778 | -97.951
3.7815 | -97.8848
3.7852 | -97.8467
3.7889 | -97.8364
3.7926 | -97.8575
3.7963 | -97.9027
3.8 | -97.9637
3.8037 | -98.0313
3.8074 | -98.1042
3.8111 | -98.1737
3.8148 | -98.2526
3.8185 | -98.3462
3.8222 | -98.449
3.8259 | -98.5708
3.8296 | -98.7201
3.8333 | -98.873
3.837 | -99.0177
3.8407 | -99.1389
3.8444 | -99.238
3.8481 | -99.2828
3.8519 | -99.278
3.8556 | -99.2088
3.8593 | -99.0958
3.863 | -98.9551
3.8667 | -98.8024
3.8704 | -98.654
3.8741 | -98.5246
3.8778 | -98.4177
3.8815 | -98.3435
3.8852 | -98.3093
3.8889 | -98.2941
3.8926 | -98.2807
3.8963 | -98.2739
3.9 | -98.2623
3.9037 | -98.2451
3.9074 | -98.2314
3.9111 | -98.234
3.9148 | -98.2523
3.9185 | -98.2957
3.9222 | -98.3436
3.9259 | -98.3924
3.9296 | -98.4373
3.9333 | -98.4739
3.937 | -98.5049
3.9407 | -98.5312
3.9444 | -98.5512
3.9481 | -98.5776
3.9519 | -98.6189
3.9556 | -98.6649
3.9593 | -98.7255
3.963 | -98.7897
3.9667 | -98.8439
3.9704 | -98.8945
3.9741 | -98.9393
3.9778 | -98.9767
3.9815 | -98.9973
3.9852 | -99.0127
3.9889 | -99.0327
3.9926 | -99.0613
3.9963 | -99.1015
4 | -99.1488
4.0037 | -99.1962
4.0074 | -99.2328
4.0111 | -99.2481
4.0148 | -99.2425
4.0185 | -99.2277
4.0222 | -99.1892
4.0259 | -99.1407
4.0296 | -99.0946
4.0333 | -99.0559
4.037 | -99.0461
4.0407 | -99.0568
4.0444 | -99.0966
4.0481 | -99.1578
4.0519 | -99.2195
4.0556 | -99.2645
4.0593 | -99.2911
4.063 | -99.2881
4.0667 | -99.2548
4.0704 | -99.2006
4.0741 | -99.1337
4.0778 | -99.0706
4.0815 | -99.0239
4.0852 | -98.9993
4.0889 | -99.0071
4.0926 | -99.03
4.0963 | -99.0548
4.1 | -99.0672
4.1037 | -99.0695
4.1074 | -99.0372
4.1111 | -98.977
4.1148 | -98.8987
4.1185 | -98.8197
4.1222 | -98.74
4.1259 | -98.6821
4.1296 | -98.6524
4.1333 | -98.6502
4.137 | -98.6763
4.1407 | -98.7311
4.1444 | -98.8161
4.1481 | -98.924
4.1519 | -99.0418
4.1556 | -99.1588
4.1593 | -99.2649
4.163 | -99.3586
4.1667 | -99.4144
4.1704 | -99.4467
4.1741 | -99.449
4.1778 | -99.4322
4.1815 | -99.3981
4.1852 | -99.3642
4.1889 | -99.3334
4.1926 | -99.3136
4.1963 | -99.3107
4.2 | -99.327
4.2037 | -99.3501
4.2074 | -99.3681
4.2111 | -99.3895
4.2148 | -99.412
4.2185 | -99.4415
4.2222 | -99.4714
4.2259 | -99.5008
4.2296 | -99.5344
4.2333 | -99.5431
4.237 | -99.5256
4.2407 | -99.4793
4.2444 | -99.4151
4.2481 | -99.3348
4.2519 | -99.2568
4.2556 | -99.1925
4.2593 | -99.1447
4.263 | -99.1141
4.2667 | -99.1134
4.2704 | -99.1269
4.2741 | -99.1751
4.2778 | -99.2195
4.2815 | -99.2745
4.2852 | -99.3455
4.2889 | -99.4372
4.2926 | -99.5381
4.2963 | -99.6385
4.3 | -99.7262
4.3037 | -99.7823
4.3074 | -99.8087
4.3111 | -99.8064
4.3148 | -99.7919
4.3185 | -99.7714
4.3222 | -99.7463
4.3259 | -99.7254
4.3296 | -99.701
4.3333 | -99.6577
4.337 | -99.6074
4.3407 | -99.5437
4.3444 | -99.4751
4.3481 | -99.4155
4.3519 | -99.3737
4.3556 | -99.3544
4.3593 | -99.3595
4.363 | -99.3894
4.3667 | -99.4348
4.3704 | -99.4832
4.3741 | -99.5229
4.3778 | -99.546
4.3815 | -99.5494
4.3852 | -99.527
4.3889 | -99.4804
4.3926 | -99.4152
4.3963 | -99.3397
4.4 | -99.2647
4.4037 | -99.2019
4.4074 | -99.1527
4.4111 | -99.1215
4.4148 | -99.1106
4.4185 | -99.1293
4.4222 | -99.1791
4.4259 | -99.2571
4.4296 | -99.3688
4.4333 | -99.4972
4.437 | -99.6307
4.4407 | -99.7653
4.4444 | -99.8749
4.4481 | -99.9544
4.4519 | -100.0031
4.4556 | -100.0309
4.4593 | -100.0348
4.463 | -100.0078
4.4667 | -99.951
4.4704 | -99.876
4.4741 | -99.7954
4.4778 | -99.7243
4.4815 | -99.6813
4.4852 | -99.663
4.4889 | -99.6571
4.4926 | -99.6611
4.4963 | -99.6692
4.5 | -99.6717
4.5037 | -99.679
4.5074 | -99.684
4.5111 | -99.6823
4.5148 | -99.6842
4.5185 | -99.6937
4.5222 | -99.709
4.5259 | -99.7198
4.5296 | -99.7321
4.5333 | -99.7543
4.537 | -99.7776
4.5407 | -99.7952
4.5444 | -99.8115
4.5481 | -99.8224
4.5519 | -99.8249
4.5556 | -99.8312
4.5593 | -99.8427
4.563 | -99.8505
4.5667 | -99.8531
4.5704 | -99.861
4.5741 | -99.8686
4.5778 | -99.8714
4.5815 | -99.8707
4.5852 | -99.8797
4.5889 | -99.888
4.5926 | -99.9064
4.5963 | -99.9361
4.6 | -99.9784
4.6037 | -100.0281
4.6074 | -100.0911
4.6111 | -100.1691
4.6148 | -100.2432
4.6185 | -100.3268
4.6222 | -100.4076
4.6259 | -100.4761
4.6296 | -100.534
4.6333 | -100.5744
4.637 | -100.5948
4.6407 | -100.6002
4.6444 | -100.5971
4.6481 | -100.5862
4.6519 | -100.5656
4.6556 | -100.5211
4.6593 | -100.4701
4.663 | -100.4115
4.6667 | -100.3456
4.6704 | -100.2736
4.6741 | -100.2061
4.6778 | -100.1442
4.6815 | -100.0959
4.6852 | -100.0696
4.6889 | -100.0587
4.6926 | -100.0631
4.6963 | -100.081
4.7 | -100.1062
4.7037 | -100.1226
4.7074 | -100.132
4.7111 | -100.1327
4.7148 | -100.1128
4.7185 | -100.0891
4.7222 | -100.0612
4.7259 | -100.0287
4.7296 | -99.9998
4.7333 | -99.9849
4.737 | -99.9909
4.7407 | -100.0122
4.7444 | -100.0374
4.7481 | -100.0719
4.7519 | -100.1116
4.7556 | -100.147
4.7593 | -100.1694
4.763 | -100.1916
4.7667 | -100.2086
4.7704 | -100.2242
4.7741 | -100.2526
4.7778 | -100.279
4.7815 | -100.3049
4.7852 | -100.3431
4.7889 | -100.376
4.7926 | -100.4152
4.7963 | -100.4644
4.8 | -100.5148
4.8037 | -100.5649
4.8074 | -100.5911
4.8111 | -100.6218
4.8148 | -100.6459
4.8185 | -100.6567
4.8222 | -100.6432
4.8259 | -100.6157
4.8296 | -100.5794
4.8333 | -100.5383
4.837 | -100.4889
4.8407 | -100.4391
4.8444 | -100.3903
4.8481 | -100.3406
4.8519 | -100.2951
4.8556 | -100.2551
4.8593 | -100.219
4.863 | -100.1834
4.8667 | -100.1469
4.8704 | -100.1262
4.8741 | -100.1189
4.8778 | -100.121
4.8815 | -100.1367
4.8852 | -100.1597
4.8889 | -100.189
4.8926 | -100.2208
4.8963 | -100.2626
4.9 | -100.3076
4.9037 | -100.3441
4.9074 | -100.3806
4.9111 | -100.4186
4.9148 | -100.4519
4.9185 | -100.4838
4.9222 | -100.5143
4.9259 | -100.5388
4.9296 | -100.5685
4.9333 | -100.6054
4.937 | -100.6478
4.9407 | -100.6986
4.9444 | -100.7533
4.9481 | -100.7981
4.9519 | -100.8436
4.9556 | -100.8887
4.9593 | -100.9258
4.963 | -100.9469
4.9667 | -100.9483
4.9704 | -100.936
4.9741 | -100.9034
4.9778 | -100.8615
4.9815 | -100.8209
4.9852 | -100.7799
4.9889 | -100.7469
4.9926 | -100.7337
4.9963 | -100.7236
5 | -100.7258
5.0037 | -100.7374
5.0074 | -100.7426
5.0111 | -100.7358
5.0148 | -100.7227
5.0185 | -100.7097
5.0222 | -100.6899
5.0259 | -100.6635
5.0296 | -100.6266
5.0333 | -100.5864
5.037 | -100.5519
5.0407 | -100.5207
5.0444 | -100.4996
5.0481 | -100.4932
5.0519 | -100.5013
5.0556 | -100.5144
5.0593 | -100.5304
5.063 | -100.5541
5.0667 | -100.5783
5.0704 | -100.5926
5.0741 | -100.6085
5.0778 | -100.62
5.0815 | -100.6286
5.0852 | -100.6345
5.0889 | -100.6454
5.0926 | -100.6535
5.0963 | -100.6567
5.1 | -100.6655
5.1037 | -100.6855
5.1074 | -100.7182
5.1111 | -100.7588
5.1148 | -100.8061
5.1185 | -100.8672
5.1222 | -100.9383
5.1259 | -101.0095
5.1296 | -101.0778
5.1333 | -101.1418
5.137 | -101.1969
5.1407 | -101.2319
5.1444 | -101.2519
5.1481 | -101.2607
5.1519 | -101.255
5.1556 | -101.2404
5.1593 | -101.2314
5.163 | -101.2342
5.1667 | -101.2366
5.1704 | -101.2496
5.1741 | -101.2662
5.1778 | -101.2788
5.1815 | -101.2819
5.1852 | -101.2676
5.1889 | -101.2322
5.1926 | -101.1814
5.1963 | -101.1198
5.2 | -101.0504
5.2037 | -100.9816
5.2074 | -100.9119
5.2111 | -100.8541
5.2148 | -100.8099
5.2185 | -100.783
5.2222 | -100.7743
5.2259 | -100.7803
5.2296 | -100.7957
5.2333 | -100.8185
5.237 | -100.8382
5.2407 | -100.8505
5.2444 | -100.8673
5.2481 | -100.8788
5.2519 | -100.8876
5.2556 | -100.8979
5.2593 | -100.9053
5.263 | -100.9162
5.2667 | -100.9348
5.2704 | -100.9462
5.2741 | -100.9585
5.2778 | -100.9802
5.2815 | -101.0138
5.2852 | -101.0556
5.2889 | -101.0943
5.2926 | -101.1358
5.2963 | -101.1821
5.3 | -101.224
5.3037 | -101.266
5.3074 | -101.3023
5.3111 | -101.3285
5.3148 | -101.3413
5.3185 | -101.344
5.3222 | -101.3295
5.3259 | -101.3064
5.3296 | -101.2787
5.3333 | -101.2474
5.337 | -101.2188
5.3407 | -101.2282
5.3444 | -101.2298
5.3481 | -101.2392
5.3519 | -101.2625
5.3556 | -101.2933
5.3593 | -101.318
5.363 | -101.3374
5.3667 | -101.3568
5.3704 | -101.3622
5.3741 | -101.3483
5.3778 | -101.3202
5.3815 | -101.2779
5.3852 | -101.2237
5.3889 | -101.1598
5.3926 | -101.0902
5.3963 | -101.025
5.4 | -100.9651
5.4037 | -100.914
5.4074 | -100.8744
5.4111 | -100.8506
5.4148 | -100.8453
5.4185 | -100.8518
5.4222 | -100.8697
5.4259 | -100.897
5.4296 | -100.9431
5.4333 | -100.999
5.437 | -101.0595
5.4407 | -101.1228
5.4444 | -101.19
5.4481 | -101.2506
5.4519 | -101.3008
5.4556 | -101.3406
5.4593 | -101.3664
5.463 | -101.3725
5.4667 | -101.3654
5.4704 | -101.3414
5.4741 | -101.3008
5.4778 | -101.2537
5.4815 | -101.2047
5.4852 | -101.1583
5.4889 | -101.1235
5.4926 | -101.1002
5.4963 | -101.0883
5.5 | -101.0862
5.5037 | -101.0941
5.5074 | -101.1061
5.5111 | -101.1197
5.5148 | -101.1229
5.5185 | -101.1306
5.5222 | -101.1451
5.5259 | -101.1548
5.5296 | -101.1642
5.5333 | -101.1769
5.537 | -101.1855
5.5407 | -101.1908
5.5444 | -101.1954
5.5481 | -101.1942
5.5519 | -101.1924
5.5556 | -101.1966
5.5593 | -101.2022
5.563 | -101.2099
5.5667 | -101.2228
5.5704 | -101.2456
5.5741 | -101.2794
5.5778 | -101.3214
5.5815 | -101.374
5.5852 | -101.4317
5.5889 | -101.4906
5.5926 | -101.5457
5.5963 | -101.5927
5.6 | -101.6343
5.6037 | -101.662
5.6074 | -101.6788
5.6111 | -101.697
5.6148 | -101.7146
5.6185 | -101.7287
5.6222 | -101.7498
5.6259 | -101.7823
5.6296 | -101.8128
5.6333 | -101.8465
5.637 | -101.8746
5.6407 | -101.8999
5.6444 | -101.925
5.6481 | -101.9435
5.6519 | -101.9412
5.6556 | -101.9234
5.6593 | -101.8926
5.663 | -101.8489
5.6667 | -101.7891
5.6704 | -101.7334
5.6741 | -101.6798
5.6778 | -101.6279
5.6815 | -101.591
5.6852 | -101.567
5.6889 | -101.5584
5.6926 | -101.5626
5.6963 | -101.5791
5.7 | -101.6157
5.7037 | -101.6739
5.7074 | -101.7399
5.7111 | -101.8085
5.7148 | -101.8825
5.7185 | -101.9598
5.7222 | -102.0217
5.7259 | -102.053
5.7296 | -102.0686
5.7333 | -102.0588
5.737 | -102.0405
5.7407 | -102.0042
5.7444 | -101.9523
5.7481 | -101.8976
5.7519 | -101.8438
5.7556 | -101.7902
5.7593 | -101.748
5.763 | -101.7167
5.7667 | -101.6936
5.7704 | -101.6802
5.7741 | -101.674
5.7778 | -101.6764
5.7815 | -101.6799
5.7852 | -101.6765
5.7889 | -101.6719
5.7926 | -101.6727
5.7963 | -101.6803
5.8 | -101.6966
5.8037 | -101.72
5.8074 | -101.7618
5.8111 | -101.8226
5.8148 | -101.892
5.8185 | -101.9693
5.8222 | -102.0486
5.8259 | -102.1237
5.8296 | -102.1923
5.8333 | -102.2447
5.837 | -102.2666
5.8407 | -102.2563
5.8444 | -102.2169
5.8481 | -102.1592
5.8519 | -102.0858
5.8556 | -101.9977
5.8593 | -101.9124
5.863 | -101.8375
5.8667 | -101.7714
5.8704 | -101.7195
5.8741 | -101.7108
5.8778 | -101.6859
5.8815 | -101.6732
5.8852 | -101.6693
5.8889 | -101.6809
5.8926 | -101.7039
5.8963 | -101.7321
5.9 | -101.7678
5.9037 | -101.8128
5.9074 | -101.8672
5.9111 | -101.919
5.9148 | -101.9635
5.9185 | -102.0007
5.9222 | -102.026
5.9259 | -102.0448
5.9296 | -102.0561
5.9333 | -102.0566
5.937 | -102.0472
5.9407 | -102.0337
5.9444 | -102.0282
5.9481 | -102.0132
5.9519 | -102.0037
5.9556 | -102.0082
5.9593 | -102.0154
5.963 | -102.0222
5.9667 | -102.0292
5.9704 | -102.037
5.9741 | -102.0411
5.9778 | -102.0429
5.9815 | -102.0343
5.9852 | -102.0178
5.9889 | -102.0029
5.9926 | -101.9811
5.9963 | -101.9465
6 | -101.9103
6.0037 | -101.8753
6.0074 | -101.8381
6.0111 | -101.7978
6.0148 | -101.7629
6.0185 | -101.7335
6.0222 | -101.7074
6.0259 | -101.6945
6.0296 | -101.6933
6.0333 | -101.7033
6.037 | -101.7261
6.0407 | -101.7681
6.0444 | -101.8122
6.0481 | -101.859
6.0519 | -101.914
6.0556 | -101.9727
6.0593 | -102.0374
6.063 | -102.0989
6.0667 | -102.1526
6.0704 | -102.1971
6.0741 | -102.2303
6.0778 | -102.2477
6.0815 | -102.2466
6.0852 | -102.2324
6.0889 | -102.207
6.0926 | -102.1708
6.0963 | -102.1268
6.1 | -102.0886
6.1037 | -102.0578
6.1074 | -102.0297
6.1111 | -102.0086
6.1148 | -101.9933
6.1185 | -101.9821
6.1222 | -101.9718
6.1259 | -101.9651
6.1296 | -101.9565
6.1333 | -101.9457
6.137 | -101.9339
6.1407 | -101.9231
6.1444 | -101.9155
6.1481 | -101.913
6.1519 | -101.9141
6.1556 | -101.925
6.1593 | -101.9425
6.163 | -101.9601
6.1667 | -101.9824
6.1704 | -102.0095
6.1741 | -102.0353
6.1778 | -102.0572
6.1815 | -102.0816
6.1852 | -102.1047
6.1889 | -102.1272
6.1926 | -102.1421
6.1963 | -102.1548
6.2 | -102.1711
6.2037 | -102.1825
6.2074 | -102.1953
6.2111 | -102.202
6.2148 | -102.201
6.2185 | -102.1902
6.2222 | -102.1783
6.2259 | -102.1722
6.2296 | -102.1711
6.2333 | -102.1792
6.237 | -102.1944
6.2407 | -102.2214
6.2444 | -102.2586
6.2481 | -102.3015
6.2519 | -102.3524
6.2556 | -102.409
6.2593 | -102.4672
6.263 | -102.5212
6.2667 | -102.5657
6.2704 | -102.6074
6.2741 | -102.6286
6.2778 | -102.6413
6.2815 | -102.6423
6.2852 | -102.625
6.2889 | -102.5887
6.2926 | -102.536
6.2963 | -102.4757
6.3 | -102.3983
6.3037 | -102.3194
6.3074 | -102.2495
6.3111 | -102.1852
6.3148 | -102.1307
6.3185 | -102.0854
6.3222 | -102.0568
6.3259 | -102.0357
6.3296 | -102.024
6.3333 | -102.0253
6.337 | -102.0396
6.3407 | -102.0672
6.3444 | -102.0999
6.3481 | -102.1419
6.3519 | -102.1869
6.3556 | -102.2348
6.3593 | -102.2887
6.363 | -102.3372
6.3667 | -102.3885
6.3704 | -102.4278
6.3741 | -102.46
6.3778 | -102.4833
6.3815 | -102.4928
6.3852 | -102.4968
6.3889 | -102.4848
6.3926 | -102.4663
6.3963 | -102.4363
6.4 | -102.3883
6.4037 | -102.3354
6.4074 | -102.2786
6.4111 | -102.2207
6.4148 | -102.1657
6.4185 | -102.1093
6.4222 | -102.0564
6.4259 | -102.0039
6.4296 | -101.9529
6.4333 | -101.9048
6.437 | -101.8598
6.4407 | -101.8246
6.4444 | -101.7962
6.4481 | -101.7721
6.4519 | -101.7575
6.4556 | -101.7461
6.4593 | -101.7439
6.463 | -101.7595
6.4667 | -101.7879
6.4704 | -101.831
6.4741 | -101.8903
6.4778 | -101.9668
6.4815 | -102.0529
6.4852 | -102.1548
6.4889 | -102.2676
6.4926 | -102.3844
6.4963 | -102.5053
6.5 | -102.6165
6.5037 | -102.7162
6.5074 | -102.8087
6.5111 | -102.8887
6.5148 | -102.9521
6.5185 | -103.0027
6.5222 | -103.0368
6.5259 | -103.0575
6.5296 | -103.0597
6.5333 | -103.0472
6.537 | -103.0254
6.5407 | -102.9986
6.5444 | -102.9731
6.5481 | -102.9576
6.5519 | -102.9414
6.5556 | -102.9271
6.5593 | -102.9117
6.563 | -102.8859
6.5667 | -102.8629
6.5704 | -102.8354
6.5741 | -102.8045
6.5778 | -102.7716
6.5815 | -102.7345
6.5852 | -102.6925
6.5889 | -102.6539
6.5926 | -102.6164
6.5963 | -102.5861
6.6 | -102.5605
6.6037 | -102.5346
6.6074 | -102.5123
6.6111 | -102.4961
6.6148 | -102.4872
6.6185 | -102.483
6.6222 | -102.481
6.6259 | -102.4778
6.6296 | -102.4779
6.6333 | -102.4692
6.637 | -102.4533
6.6407 | -102.4325
6.6444 | -102.403
6.6481 | -102.3707
6.6519 | -102.3411
6.6556 | -102.3105
6.6593 | -102.2758
6.663 | -102.2399
6.6667 | -102.2047
6.6704 | -102.1717
6.6741 | -102.1465
6.6778 | -102.1203
6.6815 | -102.1005
6.6852 | -102.0886
6.6889 | -102.0732
6.6926 | -102.0661
6.6963 | -102.0695
6.7 | -102.0837
6.7037 | -102.1036
6.7074 | -102.1314
6.7111 | -102.1651
6.7148 | -102.209
6.7185 | -102.2615
6.7222 | -102.3247
6.7259 | -102.3942
6.7296 | -102.4612
6.7333 | -102.5353
6.737 | -102.6073
6.7407 | -102.6661
6.7444 | -102.7194
6.7481 | -102.7661
6.7519 | -102.8056
6.7556 | -102.8398
6.7593 | -102.8758
6.763 | -102.9077
6.7667 | -102.9394
6.7704 | -102.9774
6.7741 | -103.0159
6.7778 | -103.0609
6.7815 | -103.1055
6.7852 | -103.1425
6.7889 | -103.1826
6.7926 | -103.2183
6.7963 | -103.2312
6.8 | -103.227
6.8037 | -103.2105
6.8074 | -103.1753
6.8111 | -103.1229
6.8148 | -103.0522
6.8185 | -102.9689
6.8222 | -102.8751
6.8259 | -102.7746
6.8296 | -102.6759
6.8333 | -102.5744
6.837 | -102.4815
6.8407 | -102.4074
6.8444 | -102.3422
6.8481 | -102.2898
6.8519 | -102.2576
6.8556 | -102.2445
6.8593 | -102.2464
6.863 | -102.2684
6.8667 | -102.3007
6.8704 | -102.3419
6.8741 | -102.3877
6.8778 | -102.4428
6.8815 | -102.5027
6.8852 | -102.5633
6.8889 | -102.617
6.8926 | -102.6651
6.8963 | -102.7174
6.9 | -102.7642
6.9037 | -102.8055
6.9074 | -102.8354
6.9111 | -102.8568
6.9148 | -102.8719
6.9185 | -102.8777
6.9222 | -102.8769
6.9259 | -102.8694
6.9296 | -102.8549
6.9333 | -102.8408
6.937 | -102.8201
6.9407 | -102.7968
6.9444 | -102.7942
6.9481 | -102.788
6.9519 | -102.7878
6.9556 | -102.7954
6.9593 | -102.8087
6.963 | -102.8298
6.9667 | -102.8577
6.9704 | -102.8934
6.9741 | -102.9297
6.9778 | -102.9746
6.9815 | -103.0154
6.9852 | -103.0563
6.9889 | -103.0972
6.9926 | -103.1385
6.9963 | -103.1728
7 | -103.1989
7.0037 | -103.2196
7.0074 | -103.2346
7.0111 | -103.2403
7.0148 | -103.2395
7.0185 | -103.2259
7.0222 | -103.2022
7.0259 | -103.1752
7.0296 | -103.1418
7.0333 | -103.0999
7.037 | -103.0551
7.0407 | -103.0128
7.0444 | -102.9721
7.0481 | -102.9366
7.0519 | -102.9086
7.0556 | -102.884
7.0593 | -102.8659
7.063 | -102.8529
7.0667 | -102.8484
7.0704 | -102.8387
7.0741 | -102.8354
7.0778 | -102.8372
7.0815 | -102.8392
7.0852 | -102.8324
7.0889 | -102.8299
7.0926 | -102.8275
7.0963 | -102.821
7.1 | -102.81
7.1037 | -102.8015
7.1074 | -102.7904
7.1111 | -102.7769
7.1148 | -102.7585
7.1185 | -102.7441
7.1222 | -102.7328
7.1259 | -102.7336
7.1296 | -102.7403
7.1333 | -102.7584
7.137 | -102.789
7.1407 | -102.8316
7.1444 | -102.8824
7.1481 | -102.9408
7.1519 | -103.0024
7.1556 | -103.0704
7.1593 | -103.1391
7.163 | -103.2064
7.1667 | -103.2648
7.1704 | -103.3133
7.1741 | -103.3536
7.1778 | -103.3941
7.1815 | -103.4211
7.1852 | -103.4374
7.1889 | -103.4515
7.1926 | -103.4694
7.1963 | -103.4783
7.2 | -103.4882
7.2037 | -103.5057
7.2074 | -103.5252
7.2111 | -103.5446
7.2148 | -103.5676
7.2185 | -103.5888
7.2222 | -103.6081
7.2259 | -103.6213
7.2296 | -103.6248
7.2333 | -103.6146
7.237 | -103.6015
7.2407 | -103.5725
7.2444 | -103.5331
7.2481 | -103.4881
7.2519 | -103.4426
7.2556 | -103.3944
7.2593 | -103.3426
7.263 | -103.2811
7.2667 | -103.2179
7.2704 | -103.159
7.2741 | -103.1026
7.2778 | -103.0481
7.2815 | -103.0016
7.2852 | -102.9623
7.2889 | -102.9324
7.2926 | -102.9123
7.2963 | -102.8963
7.3 | -102.8857
7.3037 | -102.8848
7.3074 | -102.8879
7.3111 | -102.8848
7.3148 | -102.8842
7.3185 | -102.8796
7.3222 | -102.873
7.3259 | -102.8629
7.3296 | -102.8426
7.3333 | -102.8179
7.337 | -102.7901
7.3407 | -102.7626
7.3444 | -102.7376
7.3481 | -102.7183
7.3519 | -102.7109
7.3556 | -102.7124
7.3593 | -102.7212
7.363 | -102.7447
7.3667 | -102.7844
7.3704 | -102.8419
7.3741 | -102.9084
7.3778 | -102.983
7.3815 | -103.0653
7.3852 | -103.1428
7.3889 | -103.2113
7.3926 | -103.2724
7.3963 | -103.3236
7.4 | -103.3628
7.4037 | -103.3918
7.4074 | -103.4023
7.4111 | -103.4029
7.4148 | -103.3927
7.4185 | -103.3849
7.4222 | -103.368
7.4259 | -103.3472
7.4296 | -103.3226
7.4333 | -103.3009
7.437 | -103.2845
7.4407 | -103.2695
7.4444 | -103.2598
7.4481 | -103.2511
7.4519 | -103.2444
7.4556 | -103.2457
7.4593 | -103.2487
7.463 | -103.2539
7.4667 | -103.263
7.4704 | -103.2711
7.4741 | -103.281
7.4778 | -103.2925
7.4815 | -103.2977
7.4852 | -103.3018
7.4889 | -103.2985
7.4926 | -103.2931
7.4963 | -103.2903
7.5 | -103.2844
7.5037 | -103.277
7.5074 | -103.2678
7.5111 | -103.257
7.5148 | -103.2429
7.5185 | -103.2289
7.5222 | -103.2131
7.5259 | -103.2005
7.5296 | -103.1922
7.5333 | -103.1923
7.537 | -103.2006
7.5407 | -103.2171
7.5444 | -103.2358
7.5481 | -103.2582
7.5519 | -103.2909
7.5556 | -103.3286
7.5593 | -103.3686
7.563 | -103.4112
7.5667 | -103.4523
7.5704 | -103.4878
7.5741 | -103.5201
7.5778 | -103.5453
7.5815 | -103.5606
7.5852 | -103.5675
7.5889 | -103.5602
7.5926 | -103.547
7.5963 | -103.5307
7.6 | -103.5127
7.6037 | -103.4937
7.6074 | -103.4726
7.6111 | -103.4521
7.6148 | -103.4331
7.6185 | -103.4175
7.6222 | -103.4043
7.6259 | -103.3992
7.6296 | -103.3973
7.6333 | -103.3941
7.637 | -103.3987
7.6407 | -103.4037
7.6444 | -103.4083
7.6481 | -103.4149
7.6519 | -103.4239
7.6556 | -103.4331
7.6593 | -103.4411
7.663 | -103.4461
7.6667 | -103.443
7.6704 | -103.4411
7.6741 | -103.4408
7.6778 | -103.4388
7.6815 | -103.4358
7.6852 | -103.4322
7.6889 | -103.4266
7.6926 | -103.4231
7.6963 | -103.4243
7.7 | -103.4223
7.7037 | -103.4273
7.7074 | -103.4355
7.7111 | -103.4467
7.7148 | -103.4563
7.7185 | -103.4633
7.7222 | -103.4726
7.7259 | -103.4888
7.7296 | -103.5053
7.7333 | -103.5184
7.737 | -103.5329
7.7407 | -103.5466
7.7444 | -103.5593
7.7481 | -103.5687
7.7519 | -103.5677
7.7556 | -103.5656
7.7593 | -103.5609
7.763 | -103.5551
7.7667 | -103.5495
7.7704 | -103.5449
7.7741 | -103.5445
7.7778 | -103.5481
7.7815 | -103.5537
7.7852 | -103.5611
7.7889 | -103.5696
7.7926 | -103.5792
7.7963 | -103.5898
7.8 | -103.6057
7.8037 | -103.6223
7.8074 | -103.6312
7.8111 | -103.6361
7.8148 | -103.6461
7.8185 | -103.6518
7.8222 | -103.6496
7.8259 | -103.6446
7.8296 | -103.6383
7.8333 | -103.6316
7.837 | -103.6191
7.8407 | -103.6017
7.8444 | -103.5909
7.8481 | -103.585
7.8519 | -103.5819
7.8556 | -103.5861
7.8593 | -103.6014
7.863 | -103.622
7.8667 | -103.65
7.8704 | -103.6824
7.8741 | -103.7164
7.8778 | -103.7475
7.8815 | -103.7737
7.8852 | -103.7965
7.8889 | -103.8158
7.8926 | -103.8241
7.8963 | -103.8225
7.9 | -103.8164
7.9037 | -103.8075
7.9074 | -103.7958
7.9111 | -103.7851
7.9148 | -103.7723
7.9185 | -103.76
7.9222 | -103.7527
7.9259 | -103.7496
7.9296 | -103.7471
7.9333 | -103.7511
7.937 | -103.7668
7.9407 | -103.787
7.9444 | -103.8097
7.9481 | -103.8354
7.9519 | -103.8638
7.9556 | -103.8936
7.9593 | -103.9197
7.963 | -103.9429
7.9667 | -103.9626
7.9704 | -103.9754
7.9741 | -103.9822
7.9778 | -103.988
7.9815 | -103.9881
7.9852 | -103.9832
7.9889 | -103.9813
7.9926 | -103.9772
7.9963 | -103.9674
8 | -103.9566
8.0037 | -103.9455
8.0074 | -103.9381
8.0111 | -103.9263
8.0148 | -103.9076
8.0185 | -103.89
8.0222 | -103.8734
8.0259 | -103.8572
8.0296 | -103.8367
8.0333 | -103.8154
8.037 | -103.797
8.0407 | -103.7771
8.0444 | -103.7517
8.0481 | -103.7243
8.0519 | -103.6979
8.0556 | -103.6735
8.0593 | -103.6566
8.063 | -103.6453
8.0667 | -103.6382
8.0704 | -103.6349
8.0741 | -103.6343
8.0778 | -103.6368
8.0815 | -103.6471
8.0852 | -103.6649
8.0889 | -103.6866
8.0926 | -103.7111
8.0963 | -103.7352
8.1 | -103.7582
8.1037 | -103.7814
8.1074 | -103.8011
8.1111 | -103.8198
8.1148 | -103.8362
8.1185 | -103.8446
8.1222 | -103.847
8.1259 | -103.849
8.1296 | -103.8488
8.1333 | -103.847
8.137 | -103.8454
8.1407 | -103.843
8.1444 | -103.8371
8.1481 | -103.8313
8.1519 | -103.8304
8.1556 | -103.8319
8.1593 | -103.8361
8.163 | -103.8475
8.1667 | -103.8665
8.1704 | -103.8897
8.1741 | -103.9134
8.1778 | -103.9377
8.1815 | -103.965
8.1852 | -103.9969
8.1889 | -104.028
8.1926 | -104.0594
8.1963 | -104.09
8.2 | -104.1172
8.2037 | -104.1379
8.2074 | -104.1515
8.2111 | -104.1606
8.2148 | -104.1626
8.2185 | -104.1597
8.2222 | -104.1485
8.2259 | -104.1289
8.2296 | -104.1069
8.2333 | -104.0799
8.237 | -104.0511
8.2407 | -104.0184
8.2444 | -103.9873
8.2481 | -103.9525
8.2519 | -103.9206
8.2556 | -103.8904
8.2593 | -103.8604
8.263 | -103.8323
8.2667 | -103.8106
8.2704 | -103.7922
8.2741 | -103.7747
8.2778 | -103.7612
8.2815 | -103.7525
8.2852 | -103.7474
8.2889 | -103.7472
8.2926 | -103.7557
8.2963 | -103.7693
8.3 | -103.7857
8.3037 | -103.805
8.3074 | -103.8256
8.3111 | -103.8528
8.3148 | -103.8807
8.3185 | -103.9067
8.3222 | -103.932
8.3259 | -103.9455
8.3296 | -103.958
8.3333 | -103.9664
8.337 | -103.972
8.3407 | -103.9744
8.3444 | -103.9721
8.3481 | -103.9688
8.3519 | -103.9599
8.3556 | -103.9517
8.3593 | -103.9452
8.363 | -103.9391
8.3667 | -103.9273
8.3704 | -103.9165
8.3741 | -103.9108
8.3778 | -103.9062
8.3815 | -103.9038
8.3852 | -103.9048
8.3889 | -103.9064
8.3926 | -103.9134
8.3963 | -103.9239
8.4 | -103.9385
8.4037 | -103.9483
8.4074 | -103.9582
8.4111 | -103.9698
8.4148 | -103.9799
8.4185 | -103.9907
8.4222 | -103.9967
8.4259 | -103.9968
8.4296 | -103.9968
8.4333 | -103.9873
8.437 | -103.9711
8.4407 | -103.9524
8.4444 | -103.9293
8.4481 | -103.9093
8.4519 | -103.8856
8.4556 | -103.8637
8.4593 | -103.8438
8.463 | -103.8281
8.4667 | -103.8188
8.4704 | -103.8156
8.4741 | -103.822
8.4778 | -103.8389
8.4815 | -103.866
8.4852 | -103.9013
8.4889 | -103.9463
8.4926 | -103.9971
8.4963 | -104.0511
8.5 | -104.1096
8.5037 | -104.1656
8.5074 | -104.2212
8.5111 | -104.2756
8.5148 | -104.3181
8.5185 | -104.3545
8.5222 | -104.3767
8.5259 | -104.3892
8.5296 | -104.392
8.5333 | -104.3802
8.537 | -104.3577
8.5407 | -104.3315
8.5444 | -104.2997
8.5481 | -104.2635
8.5519 | -104.2363
8.5556 | -104.2069
8.5593 | -104.183
8.563 | -104.1701
8.5667 | -104.1644
8.5704 | -104.1638
8.5741 | -104.1714
8.5778 | -104.1845
8.5815 | -104.2054
8.5852 | -104.233
8.5889 | -104.2655
8.5926 | -104.3042
8.5963 | -104.3446
8.6 | -104.3812
8.6037 | -104.417
8.6074 | -104.4499
8.6111 | -104.4811
8.6148 | -104.5082
8.6185 | -104.5262
8.6222 | -104.5381
8.6259 | -104.5407
8.6296 | -104.5394
8.6333 | -104.5296
8.637 | -104.5115
8.6407 | -104.4906
8.6444 | -104.463
8.6481 | -104.4327
8.6519 | -104.3994
8.6556 | -104.3594
8.6593 | -104.3186
8.663 | -104.2787
8.6667 | -104.2374
8.6704 | -104.2041
8.6741 | -104.1761
8.6778 | -104.1458
8.6815 | -104.1237
8.6852 | -104.1117
8.6889 | -104.1083
8.6926 | -104.1075
8.6963 | -104.1162
8.7 | -104.1325
8.7037 | -104.1535
8.7074 | -104.1796
8.7111 | -104.2064
8.7148 | -104.2367
8.7185 | -104.2672
8.7222 | -104.2958
8.7259 | -104.3234
8.7296 | -104.3521
8.7333 | -104.3789
8.737 | -104.4046
8.7407 | -104.4298
8.7444 | -104.4505
8.7481 | -104.4694
8.7519 | -104.4889
8.7556 | -104.5081
8.7593 | -104.5254
8.763 | -104.54
8.7667 | -104.553
8.7704 | -104.5639
8.7741 | -104.576
8.7778 | -104.5888
8.7815 | -104.6006
8.7852 | -104.6092
8.7889 | -104.6165
8.7926 | -104.6211
8.7963 | -104.6254
8.8 | -104.6269
8.8037 | -104.6309
8.8074 | -104.6367
8.8111 | -104.6428
8.8148 | -104.6425
8.8185 | -104.641
8.8222 | -104.6402
8.8259 | -104.6333
8.8296 | -104.628
8.8333 | -104.6231
8.837 | -104.6194
8.8407 | -104.6123
8.8444 | -104.5979
8.8481 | -104.576
8.8519 | -104.554
8.8556 | -104.5311
8.8593 | -104.5022
8.863 | -104.4678
8.8667 | -104.4319
8.8704 | -104.3927
8.8741 | -104.3541
8.8778 | -104.3183
8.8815 | -104.2757
8.8852 | -104.2315
8.8889 | -104.1942
8.8926 | -104.1599
8.8963 | -104.1312
8.9 | -104.1053
8.9037 | -104.0846
8.9074 | -104.0719
8.9111 | -104.0685
8.9148 | -104.0722
8.9185 | -104.0808
8.9222 | -104.0952
8.9259 | -104.1134
8.9296 | -104.1366
8.9333 | -104.1588
8.937 | -104.181
8.9407 | -104.2029
8.9444 | -104.23
8.9481 | -104.2572
8.9519 | -104.2815
8.9556 | -104.3046
8.9593 | -104.3213
8.963 | -104.3357
8.9667 | -104.3488
8.9704 | -104.3648
8.9741 | -104.382
8.9778 | -104.4002
8.9815 | -104.4128
8.9852 | -104.4217
8.9889 | -104.4344
8.9926 | -104.4446
8.9963 | -104.4532
9 | -104.4668
9.0037 | -104.481
9.0074 | -104.4964
9.0111 | -104.5151
9.0148 | -104.5332
9.0185 | -104.5435
9.0222 | -104.5574
9.0259 | -104.5701
9.0296 | -104.587
9.0333 | -104.6004
9.037 | -104.6036
9.0407 | -104.6071
9.0444 | -104.6109
9.0481 | -104.6089
9.0519 | -104.6046
9.0556 | -104.5962
9.0593 | -104.581
9.063 | -104.5622
9.0667 | -104.5402
9.0704 | -104.5113
9.0741 | -104.4775
9.0778 | -104.443
9.0815 | -104.3914
9.0852 | -104.3527
9.0889 | -104.3152
9.0926 | -104.2799
9.0963 | -104.244
9.1 | -104.2111
9.1037 | -104.1818
9.1074 | -104.16
9.1111 | -104.1475
9.1148 | -104.1462
9.1185 | -104.1544
9.1222 | -104.1715
9.1259 | -104.1991
9.1296 | -104.2345
9.1333 | -104.2743
9.137 | -104.3221
9.1407 | -104.3718
9.1444 | -104.4273
9.1481 | -104.4881
9.1519 | -104.5461
9.1556 | -104.597
9.1593 | -104.6386
9.163 | -104.667
9.1667 | -104.6885
9.1704 | -104.7012
9.1741 | -104.6946
9.1778 | -104.6789
9.1815 | -104.6531
9.1852 | -104.6195
9.1889 | -104.5821
9.1926 | -104.5413
9.1963 | -104.4988
9.2 | -104.4558
9.2037 | -104.4153
9.2074 | -104.3785
9.2111 | -104.3464
9.2148 | -104.3223
9.2185 | -104.3032
9.2222 | -104.288
9.2259 | -104.2782
9.2296 | -104.2748
9.2333 | -104.2772
9.237 | -104.2856
9.2407 | -104.2987
9.2444 | -104.3106
9.2481 | -104.3233
9.2519 | -104.3366
9.2556 | -104.3539
9.2593 | -104.3732
9.263 | -104.393
9.2667 | -104.4146
9.2704 | -104.4366
9.2741 | -104.4537
9.2778 | -104.471
9.2815 | -104.489
9.2852 | -104.5018
9.2889 | -104.5092
9.2926 | -104.5096
9.2963 | -104.5131
9.3 | -104.5185
9.3037 | -104.5188
9.3074 | -104.5177
9.3111 | -104.5147
9.3148 | -104.5044
9.3185 | -104.4939
9.3222 | -104.484
9.3259 | -104.4682
9.3296 | -104.4536
9.3333 | -104.4428
9.337 | -104.4335
9.3407 | -104.4232
9.3444 | -104.4111
9.3481 | -104.4034
9.3519 | -104.3993
9.3556 | -104.4006
9.3593 | -104.402
9.363 | -104.404
9.3667 | -104.4076
9.3704 | -104.415
9.3741 | -104.4173
9.3778 | -104.4182
9.3815 | -104.4226
9.3852 | -104.4258
9.3889 | -104.4225
9.3926 | -104.4211
9.3963 | -104.417
9.4 | -104.4124
9.4037 | -104.4063
9.4074 | -104.3998
9.4111 | -104.3952
9.4148 | -104.388
9.4185 | -104.38
9.4222 | -104.3774
9.4259 | -104.3804
9.4296 | -104.3871
9.4333 | -104.3874
9.437 | -104.3914
9.4407 | -104.3993
9.4444 | -104.4099
9.4481 | -104.4285
9.4519 | -104.4512
9.4556 | -104.4767
9.4593 | -104.5038
9.463 | -104.5353
9.4667 | -104.5685
9.4704 | -104.6024
9.4741 | -104.637
9.4778 | -104.6732
9.4815 | -104.7099
9.4852 | -104.7406
9.4889 | -104.7685
9.4926 | -104.7865
9.4963 | -104.8015
9.5 | -104.8121
9.5037 | -104.8171
9.5074 | -104.815
9.5111 | -104.8087
9.5148 | -104.7977
9.5185 | -104.7817
9.5222 | -104.7637
9.5259 | -104.7453
9.5296 | -104.7208
9.5333 | -104.6941
9.537 | -104.67
9.5407 | -104.6468
9.5444 | -104.6235
9.5481 | -104.6026
9.5519 | -104.587
9.5556 | -104.5709
9.5593 | -104.5564
9.563 | -104.5449
9.5667 | -104.5337
9.5704 | -104.5278
9.5741 | -104.5203
9.5778 | -104.5188
9.5815 | -104.5135
9.5852 | -104.5066
9.5889 | -104.4993
9.5926 | -104.4945
9.5963 | -104.4857
9.6 | -104.4756
9.6037 | -104.4636
9.6074 | -104.4472
9.6111 | -104.4308
9.6148 | -104.4208
9.6185 | -104.4061
9.6222 | -104.3928
9.6259 | -104.3843
9.6296 | -104.3776
9.6333 | -104.3745
9.637 | -104.3754
9.6407 | -104.3762
9.6444 | -104.3798
9.6481 | -104.3882
9.6519 | -104.4022
9.6556 | -104.416
9.6593 | -104.4301
9.663 | -104.4454
9.6667 | -104.4646
9.6704 | -104.4831
9.6741 | -104.5039
9.6778 | -104.5213
9.6815 | -104.5394
9.6852 | -104.5601
9.6889 | -104.5816
9.6926 | -104.6011
9.6963 | -104.6163
9.7 | -104.6334
9.7037 | -104.6519
9.7074 | -104.6659
9.7111 | -104.681
9.7148 | -104.6933
9.7185 | -104.6967
9.7222 | -104.707
9.7259 | -104.7126
9.7296 | -104.7144
9.7333 | -104.7066
9.737 | -104.6943
9.7407 | -104.6791
9.7444 | -104.6608
9.7481 | -104.6421
9.7519 | -104.6191
9.7556 | -104.5902
9.7593 | -104.5637
9.763 | -104.5396
9.7667 | -104.5167
9.7704 | -104.4955
9.7741 | -104.4816
9.7778 | -104.4735
9.7815 | -104.4671
9.7852 | -104.4685
9.7889 | -104.4765
9.7926 | -104.4925
9.7963 | -104.5123
9.8 | -104.5382
9.8037 | -104.5688
9.8074 | -104.6038
9.8111 | -104.6387
9.8148 | -104.679
9.8185 | -104.7205
9.8222 | -104.7622
9.8259 | -104.7969
9.8296 | -104.8277
9.8333 | -104.8547
9.837 | -104.8756
9.8407 | -104.8936
9.8444 | -104.904
9.8481 | -104.9104
9.8519 | -104.915
9.8556 | -104.913
9.8593 | -104.9051
9.863 | -104.8991
9.8667 | -104.8891
9.8704 | -104.8784
9.8741 | -104.8698
9.8778 | -104.8654
9.8815 | -104.8643
9.8852 | -104.866
9.8889 | -104.8657
9.8926 | -104.8647
9.8963 | -104.8711
9.9 | -104.8842
9.9037 | -104.8999
9.9074 | -104.9175
9.9111 | -104.9334
9.9148 | -104.949
9.9185 | -104.9627
9.9222 | -104.9768
9.9259 | -104.9936
9.9296 | -105.0134
9.9333 | -105.0292
9.937 | -105.0448
9.9407 | -105.0585
9.9444 | -105.0681
9.9481 | -105.0754
9.9519 | -105.0843
9.9556 | -105.0924
9.9593 | -105.0937
9.963 | -105.0962
9.9667 | -105.0997
9.9704 | -105.1067
9.9741 | -105.1117
9.9778 | -105.1107
9.9815 | -105.1079
9.9852 | -105.1045
9.9889 | -105.0979
9.9926 | -105.0859
9.9963 | -105.076
10 | -105.0646
10.0037 | -105.0533
10.0074 | -105.0362
10.0111 | -105.0173
10.0148 | -105.0011
10.0185 | -104.9832
10.0222 | -104.9634
10.0259 | -104.9471
10.0296 | -104.9306
10.0333 | -104.9148
10.037 | -104.9038
10.0407 | -104.8915
10.0444 | -104.8809
10.0481 | -104.8742
10.0519 | -104.8712
10.0556 | -104.8698
10.0593 | -104.8752
10.063 | -104.8814
10.0667 | -104.8868
10.0704 | -104.8939
10.0741 | -104.8943
10.0778 | -104.8972
10.0815 | -104.9024
10.0852 | -104.9129
10.0889 | -104.9233
10.0926 | -104.9275
10.0963 | -104.9275
10.1 | -104.9303
10.1037 | -104.9341
10.1074 | -104.9269
10.1111 | -104.9182
10.1148 | -104.9066
10.1185 | -104.8947
10.1222 | -104.8858
10.1259 | -104.8732
10.1296 | -104.8552
10.1333 | -104.8347
10.137 | -104.8126
10.1407 | -104.787
10.1444 | -104.7654
10.1481 | -104.7463
10.1519 | -104.7273
10.1556 | -104.7102
10.1593 | -104.6939
10.163 | -104.6777
10.1667 | -104.6611
10.1704 | -104.6497
10.1741 | -104.6385
10.1778 | -104.629
10.1815 | -104.6235
10.1852 | -104.6201
10.1889 | -104.6202
10.1926 | -104.6229
10.1963 | -104.6271
10.2 | -104.636
10.2037 | -104.6501
10.2074 | -104.6631
10.2111 | -104.6793
10.2148 | -104.6972
10.2185 | -104.7123
10.2222 | -104.7312
10.2259 | -104.7473
10.2296 | -104.7553
10.2333 | -104.7617
10.237 | -104.7698
10.2407 | -104.778
10.2444 | -104.7827
10.2481 | -104.7894
10.2519 | -104.7974
10.2556 | -104.8095
10.2593 | -104.8226
10.263 | -104.8357
10.2667 | -104.8464
10.2704 | -104.8644
10.2741 | -104.889
10.2778 | -104.9151
10.2815 | -104.9413
10.2852 | -104.9673
10.2889 | -104.9922
10.2926 | -105.0101
10.2963 | -105.0276
10.3 | -105.0445
10.3037 | -105.0597
10.3074 | -105.0695
10.3111 | -105.0708
10.3148 | -105.0683
10.3185 | -105.0696
10.3222 | -105.0669
10.3259 | -105.0645
10.3296 | -105.0606
10.3333 | -105.0556
10.337 | -105.0502
10.3407 | -105.0392
10.3444 | -105.0292
10.3481 | -105.021
10.3519 | -105.0102
10.3556 | -105.0007
10.3593 | -104.9922
10.363 | -104.9845
10.3667 | -104.9786
10.3704 | -104.9757
10.3741 | -104.9716
10.3778 | -104.9694
10.3815 | -104.9712
10.3852 | -104.9738
10.3889 | -104.9768
10.3926 | -104.9775
10.3963 | -104.9827
10.4 | -104.9883
10.4037 | -104.9962
10.4074 | -105.0089
10.4111 | -105.0232
10.4148 | -105.0381
10.4185 | -105.0517
10.4222 | -105.0608
10.4259 | -105.0702
10.4296 | -105.0801
10.4333 | -105.0878
10.437 | -105.095
10.4407 | -105.0982
10.4444 | -105.0998
10.4481 | -105.0975
10.4519 | -105.09
10.4556 | -105.076
10.4593 | -105.0581
10.463 | -105.0389
10.4667 | -105.0175
10.4704 | -104.9902
10.4741 | -104.9611
10.4778 | -104.9285
10.4815 | -104.9003
10.4852 | -104.8731
10.4889 | -104.8458
10.4926 | -104.8276
10.4963 | -104.8117
10.5 | -104.8009
10.5037 | -104.7977
10.5074 | -104.7969
10.5111 | -104.7996
10.5148 | -104.808
10.5185 | -104.8154
10.5222 | -104.8237
10.5259 | -104.8387
10.5296 | -104.8575
10.5333 | -104.8753
10.537 | -104.8926
10.5407 | -104.9109
10.5444 | -104.931
10.5481 | -104.9469
10.5519 | -104.9586
10.5556 | -104.9687
10.5593 | -104.978
10.563 | -104.9881
10.5667 | -104.998
10.5704 | -105.0128
10.5741 | -105.0227
10.5778 | -105.0319
10.5815 | -105.0414
10.5852 | -105.0505
10.5889 | -105.0589
10.5926 | -105.0651
10.5963 | -105.0722
10.6 | -105.0799
10.6037 | -105.0876
10.6074 | -105.0982
10.6111 | -105.109
10.6148 | -105.1225
10.6185 | -105.134
10.6222 | -105.146
10.6259 | -105.1614
10.6296 | -105.1812
10.6333 | -105.2016
10.637 | -105.2224
10.6407 | -105.2417
10.6444 | -105.2583
10.6481 | -105.2736
10.6519 | -105.2848
10.6556 | -105.2922
10.6593 | -105.3004
10.663 | -105.3076
10.6667 | -105.3089
10.6704 | -105.3108
10.6741 | -105.3106
10.6778 | -105.3039
10.6815 | -105.2927
10.6852 | -105.2795
10.6889 | -105.2724
10.6926 | -105.2661
10.6963 | -105.2589
10.7 | -105.2512
10.7037 | -105.2432
10.7074 | -105.2365
10.7111 | -105.2301
10.7148 | -105.2236
10.7185 | -105.22
10.7222 | -105.22
10.7259 | -105.2223
10.7296 | -105.2257
10.7333 | -105.227
10.737 | -105.2348
10.7407 | -105.2477
10.7444 | -105.2587
10.7481 | -105.2706
10.7519 | -105.2832
10.7556 | -105.295
10.7593 | -105.3051
10.763 | -105.3163
10.7667 | -105.3298
10.7704 | -105.3387
10.7741 | -105.3457
10.7778 | -105.3512
10.7815 | -105.3507
10.7852 | -105.347
10.7889 | -105.3384
10.7926 | -105.3298
10.7963 | -105.3208
10.8 | -105.3067
10.8037 | -105.2866
10.8074 | -105.2629
10.8111 | -105.2367
10.8148 | -105.2088
10.8185 | -105.1786
10.8222 | -105.15
10.8259 | -105.1197
10.8296 | -105.0869
10.8333 | -105.0567
10.837 | -105.0279
10.8407 | -104.9986
10.8444 | -104.9723
10.8481 | -104.9444
10.8519 | -104.9211
10.8556 | -104.9017
10.8593 | -104.8873
10.863 | -104.8765
10.8667 | -104.869
10.8704 | -104.8647
10.8741 | -104.8617
10.8778 | -104.8609
10.8815 | -104.864
10.8852 | -104.8657
10.8889 | -104.8733
10.8926 | -104.8888
10.8963 | -104.9026
10.9 | -104.9151
10.9037 | -104.932
10.9074 | -104.954
10.9111 | -104.9755
10.9148 | -104.9964
10.9185 | -105.0207
10.9222 | -105.0458
10.9259 | -105.0696
10.9296 | -105.0921
10.9333 | -105.1168
10.937 | -105.1404
10.9407 | -105.1657
10.9444 | -105.1892
10.9481 | -105.2119
10.9519 | -105.2334
10.9556 | -105.2561
10.9593 | -105.2761
10.963 | -105.2956
10.9667 | -105.3108
10.9704 | -105.3224
10.9741 | -105.3306
10.9778 | -105.3404
10.9815 | -105.3471
10.9852 | -105.351
10.9889 | -105.3527
10.9926 | -105.3515
10.9963 | -105.3497
11 | -105.3464
11.0037 | -105.3409
11.0074 | -105.3345
11.0111 | -105.3256
11.0148 | -105.3085
11.0185 | -105.2907
11.0222 | -105.2726
11.0259 | -105.2526
11.0296 | -105.2347
11.0333 | -105.2186
11.037 | -105.2003
11.0407 | -105.182
11.0444 | -105.166
11.0481 | -105.1509
11.0519 | -105.1372
11.0556 | -105.1279
11.0593 | -105.1225
11.063 | -105.1198
11.0667 | -105.1169
11.0704 | -105.1181
11.0741 | -105.1287
11.0778 | -105.1409
11.0815 | -105.1521
11.0852 | -105.1647
11.0889 | -105.1864
11.0926 | -105.211
11.0963 | -105.2378
11.1 | -105.2672
11.1037 | -105.2973
11.1074 | -105.3294
11.1111 | -105.3631
11.1148 | -105.3929
11.1185 | -105.4228
11.1222 | -105.455
11.1259 | -105.4876
11.1296 | -105.5195
11.1333 | -105.5473
11.137 | -105.5678
11.1407 | -105.5877
11.1444 | -105.605
11.1481 | -105.6217
11.1519 | -105.6316
11.1556 | -105.6338
11.1593 | -105.637
11.163 | -105.6414
11.1667 | -105.6444
11.1704 | -105.6432
11.1741 | -105.6402
11.1778 | -105.6375
11.1815 | -105.634
11.1852 | -105.6295
11.1889 | -105.62
11.1926 | -105.6089
11.1963 | -105.5998
11.2 | -105.5943
11.2037 | -105.5903
11.2074 | -105.5819
11.2111 | -105.5721
11.2148 | -105.5632
11.2185 | -105.5652
11.2222 | -105.5633
11.2259 | -105.5645
11.2296 | -105.5719
11.2333 | -105.58
11.237 | -105.5865
11.2407 | -105.5981
11.2444 | -105.6132
11.2481 | -105.629
11.2519 | -105.6379
11.2556 | -105.6494
11.2593 | -105.6619
11.263 | -105.6697
11.2667 | -105.6752
11.2704 | -105.6789
11.2741 | -105.678
11.2778 | -105.6736
11.2815 | -105.6692
11.2852 | -105.6624
11.2889 | -105.6506
11.2926 | -105.6338
11.2963 | -105.6144
11.3 | -105.5953
11.3037 | -105.5769
11.3074 | -105.5559
11.3111 | -105.5306
11.3148 | -105.5085
11.3185 | -105.4885
11.3222 | -105.4648
11.3259 | -105.4388
11.3296 | -105.413
11.3333 | -105.3911
11.337 | -105.3732
11.3407 | -105.3536
11.3444 | -105.3309
11.3481 | -105.3108
11.3519 | -105.297
11.3556 | -105.2893
11.3593 | -105.2829
11.363 | -105.2787
11.3667 | -105.2767
11.3704 | -105.281
11.3741 | -105.2868
11.3778 | -105.2917
11.3815 | -105.299
11.3852 | -105.3069
11.3889 | -105.313
11.3926 | -105.3206
11.3963 | -105.3288
11.4 | -105.3354
11.4037 | -105.3402
11.4074 | -105.343
11.4111 | -105.3485
11.4148 | -105.3567
11.4185 | -105.3644
11.4222 | -105.3653
11.4259 | -105.3634
11.4296 | -105.3665
11.4333 | -105.3711
11.437 | -105.3709
11.4407 | -105.3693
11.4444 | -105.372
11.4481 | -105.3744
11.4519 | -105.3742
11.4556 | -105.3764
11.4593 | -105.3807
11.463 | -105.3875
11.4667 | -105.3975
11.4704 | -105.4061
11.4741 | -105.4193
11.4778 | -105.4328
11.4815 | -105.4444
11.4852 | -105.461
11.4889 | -105.4759
11.4926 | -105.4907
11.4963 | -105.5045
11.5 | -105.5184
11.5037 | -105.5364
11.5074 | -105.5519
11.5111 | -105.5674
11.5148 | -105.5826
11.5185 | -105.5919
11.5222 | -105.5962
11.5259 | -105.6021
11.5296 | -105.6093
11.5333 | -105.6173
11.537 | -105.624
11.5407 | -105.6261
11.5444 | -105.6284
11.5481 | -105.6241
11.5519 | -105.6205
11.5556 | -105.6172
11.5593 | -105.6044
11.563 | -105.5902
11.5667 | -105.5767
11.5704 | -105.5621
11.5741 | -105.546
11.5778 | -105.5223
11.5815 | -105.4958
11.5852 | -105.4678
11.5889 | -105.4418
11.5926 | -105.4149
11.5963 | -105.3811
11.6 | -105.3466
11.6037 | -105.3183
11.6074 | -105.2886
11.6111 | -105.2582
11.6148 | -105.232
11.6185 | -105.2053
11.6222 | -105.183
11.6259 | -105.164
11.6296 | -105.1481
11.6333 | -105.1343
11.637 | -105.125
11.6407 | -105.1216
11.6444 | -105.126
11.6481 | -105.1374
11.6519 | -105.1508
11.6556 | -105.1647
11.6593 | -105.1776
11.663 | -105.1955
11.6667 | -105.2146
11.6704 | -105.2323
11.6741 | -105.2515
11.6778 | -105.2703
11.6815 | -105.2865
11.6852 | -105.3025
11.6889 | -105.3189
11.6926 | -105.3287
11.6963 | -105.3307
11.7 | -105.3328
11.7037 | -105.3365
11.7074 | -105.3401
11.7111 | -105.3401
11.7148 | -105.3445
11.7185 | -105.3469
11.7222 | -105.3485
11.7259 | -105.3544
11.7296 | -105.3565
11.7333 | -105.3623
11.737 | -105.3708
11.7407 | -105.3777
11.7444 | -105.3918
11.7481 | -105.4067
11.7519 | -105.4329
11.7556 | -105.452
11.7593 | -105.4704
11.763 | -105.4876
11.7667 | -105.5076
11.7704 | -105.5289
11.7741 | -105.5477
11.7778 | -105.5658
11.7815 | -105.5797
11.7852 | -105.5901
11.7889 | -105.605
11.7926 | -105.6147
11.7963 | -105.6184
11.8 | -105.6196
11.8037 | -105.618
11.8074 | -105.6132
11.8111 | -105.6012
11.8148 | -105.5869
11.8185 | -105.5724
11.8222 | -105.5557
11.8259 | -105.541
11.8296 | -105.5289
11.8333 | -105.5133
11.837 | -105.4988
11.8407 | -105.4836
11.8444 | -105.4688
11.8481 | -105.4558
11.8519 | -105.4451
11.8556 | -105.4369
11.8593 | -105.429
11.863 | -105.4249
11.8667 | -105.4232
11.8704 | -105.4207
11.8741 | -105.4197
11.8778 | -105.4243
11.8815 | -105.4309
11.8852 | -105.4354
11.8889 | -105.4418
11.8926 | -105.4518
11.8963 | -105.4626
11.9 | -105.4751
11.9037 | -105.4858
11.9074 | -105.4964
11.9111 | -105.5105
11.9148 | -105.5156
11.9185 | -105.5193
11.9222 | -105.5253
11.9259 | -105.5292
11.9296 | -105.5329
11.9333 | -105.5339
11.937 | -105.536
11.9407 | -105.5323
11.9444 | -105.5246
11.9481 | -105.5172
11.9519 | -105.5092
11.9556 | -105.5007
11.9593 | -105.4894
11.963 | -105.4737
11.9667 | -105.4623
11.9704 | -105.4531
11.9741 | -105.4393
11.9778 | -105.4248
11.9815 | -105.411
11.9852 | -105.3987
11.9889 | -105.3843
11.9926 | -105.3675
11.9963 | -105.3497
12 | -105.3333
12.0037 | -105.3177
12.0074 | -105.3016
12.0111 | -105.2849
12.0148 | -105.2694
12.0185 | -105.2586
12.0222 | -105.2489
12.0259 | -105.2384
12.0296 | -105.2272
12.0333 | -105.2185
12.037 | -105.2132
12.0407 | -105.2086
12.0444 | -105.2048
12.0481 | -105.2013
12.0519 | -105.1988
12.0556 | -105.198
12.0593 | -105.1992
12.063 | -105.2015
12.0667 | -105.2034
12.0704 | -105.2112
12.0741 | -105.2188
12.0778 | -105.2259
12.0815 | -105.2263
12.0852 | -105.2263
12.0889 | -105.2363
12.0926 | -105.2436
12.0963 | -105.2483
12.1 | -105.251
12.1037 | -105.2519
12.1074 | -105.251
12.1111 | -105.2475
12.1148 | -105.2442
12.1185 | -105.2407
12.1222 | -105.2314
12.1259 | -105.2199
12.1296 | -105.2071
12.1333 | -105.1918
12.137 | -105.1745
12.1407 | -105.1556
12.1444 | -105.1372
12.1481 | -105.119
12.1519 | -105.1024
12.1556 | -105.0856
12.1593 | -105.0683
12.163 | -105.0546
12.1667 | -105.0459
12.1704 | -105.0398
12.1741 | -105.0353
12.1778 | -105.0344
12.1815 | -105.0382
12.1852 | -105.0433
12.1889 | -105.0518
12.1926 | -105.062
12.1963 | -105.0727
12.2 | -105.087
12.2037 | -105.1033
12.2074 | -105.1218
12.2111 | -105.14
12.2148 | -105.1617
12.2185 | -105.1831
12.2222 | -105.2042
12.2259 | -105.2242
12.2296 | -105.2453
12.2333 | -105.2652
12.237 | -105.2802
12.2407 | -105.2978
12.2444 | -105.3138
12.2481 | -105.3309
12.2519 | -105.3499
12.2556 | -105.3731
12.2593 | -105.3959
12.263 | -105.4112
12.2667 | -105.4286
12.2704 | -105.4435
12.2741 | -105.4617
12.2778 | -105.481
12.2815 | -105.5051
12.2852 | -105.5298
12.2889 | -105.5499
12.2926 | -105.5681
12.2963 | -105.583
12.3 | -105.6012
12.3037 | -105.6192
12.3074 | -105.6328
12.3111 | -105.6447
12.3148 | -105.654
12.3185 | -105.66
12.3222 | -105.6644
12.3259 | -105.6701
12.3296 | -105.6719
12.3333 | -105.6716
12.337 | -105.6727
12.3407 | -105.6687
12.3444 | -105.6581
12.3481 | -105.6525
12.3519 | -105.6474
12.3556 | -105.6402
12.3593 | -105.6331
12.363 | -105.6289
12.3667 | -105.626
12.3704 | -105.6221
12.3741 | -105.6231
12.3778 | -105.6236
12.3815 | -105.6255
12.3852 | -105.6282
12.3889 | -105.6337
12.3926 | -105.6398
12.3963 | -105.6463
12.4 | -105.6556
12.4037 | -105.6639
12.4074 | -105.6707
12.4111 | -105.679
12.4148 | -105.6906
12.4185 | -105.6984
12.4222 | -105.7052
12.4259 | -105.7101
12.4296 | -105.7128
12.4333 | -105.7176
12.437 | -105.7154
12.4407 | -105.7097
12.4444 | -105.7029
12.4481 | -105.6961
12.4519 | -105.6843
12.4556 | -105.6716
12.4593 | -105.6543
12.463 | -105.637
12.4667 | -105.6203
12.4704 | -105.602
12.4741 | -105.5816
12.4778 | -105.5608
12.4815 | -105.5382
12.4852 | -105.5204
12.4889 | -105.5076
12.4926 | -105.4953
12.4963 | -105.4829
12.5 | -105.4714
12.5037 | -105.463
12.5074 | -105.4581
12.5111 | -105.4577
12.5148 | -105.4625
12.5185 | -105.4672
12.5222 | -105.4737
12.5259 | -105.4862
12.5296 | -105.4956
12.5333 | -105.5055
12.537 | -105.5152
12.5407 | -105.5243
12.5444 | -105.5342
12.5481 | -105.543
12.5519 | -105.5487
12.5556 | -105.5513
12.5593 | -105.5536
12.563 | -105.5559
12.5667 | -105.5569
12.5704 | -105.5553
12.5741 | -105.5507
12.5778 | -105.5452
12.5815 | -105.5398
12.5852 | -105.5383
12.5889 | -105.5352
12.5926 | -105.528
12.5963 | -105.5207
12.6 | -105.5166
12.6037 | -105.5115
12.6074 | -105.5081
12.6111 | -105.5036
12.6148 | -105.4965
12.6185 | -105.495
12.6222 | -105.4973
12.6259 | -105.5003
12.6296 | -105.5005
12.6333 | -105.5023
12.637 | -105.5048
12.6407 | -105.5074
12.6444 | -105.5078
12.6481 | -105.5092
12.6519 | -105.5145
12.6556 | -105.5227
12.6593 | -105.5317
12.663 | -105.5411
12.6667 | -105.5429
12.6704 | -105.5407
12.6741 | -105.5432
12.6778 | -105.5475
12.6815 | -105.5496
12.6852 | -105.5503
12.6889 | -105.5479
12.6926 | -105.5444
12.6963 | -105.5401
12.7 | -105.5332
12.7037 | -105.5239
12.7074 | -105.513
12.7111 | -105.5006
12.7148 | -105.4885
12.7185 | -105.4782
12.7222 | -105.4703
12.7259 | -105.4628
12.7296 | -105.4584
12.7333 | -105.4572
12.737 | -105.4558
12.7407 | -105.4545
12.7444 | -105.4546
12.7481 | -105.4586
12.7519 | -105.4662
12.7556 | -105.4778
12.7593 | -105.4921
12.763 | -105.5076
12.7667 | -105.5242
12.7704 | -105.5423
12.7741 | -105.5632
12.7778 | -105.5849
12.7815 | -105.6043
12.7852 | -105.6254
12.7889 | -105.6464
12.7926 | -105.6697
12.7963 | -105.695
12.8 | -105.7172
12.8037 | -105.7349
12.8074 | -105.7514
12.8111 | -105.7656
12.8148 | -105.7769
12.8185 | -105.785
12.8222 | -105.784
12.8259 | -105.789
12.8296 | -105.7933
12.8333 | -105.797
12.837 | -105.7974
12.8407 | -105.7922
12.8444 | -105.7887
12.8481 | -105.7836
12.8519 | -105.7813
12.8556 | -105.7824
12.8593 | -105.7798
12.863 | -105.7765
12.8667 | -105.7733
12.8704 | -105.7686
12.8741 | -105.7686
12.8778 | -105.7675
12.8815 | -105.7651
12.8852 | -105.7663
12.8889 | -105.7721
12.8926 | -105.778
12.8963 | -105.7874
12.9 | -105.7944
12.9037 | -105.7998
12.9074 | -105.8033
12.9111 | -105.8087
12.9148 | -105.8156
12.9185 | -105.8231
12.9222 | -105.8347
12.9259 | -105.8471
12.9296 | -105.8557
12.9333 | -105.8644
12.937 | -105.877
12.9407 | -105.8893
12.9444 | -105.8996
12.9481 | -105.9087
12.9519 | -105.9171
12.9556 | -105.9278
12.9593 | -105.9389
12.963 | -105.9463
12.9667 | -105.9524
12.9704 | -105.956
12.9741 | -105.9579
12.9778 | -105.9567
12.9815 | -105.9521
12.9852 | -105.9472
12.9889 | -105.9441
12.9926 | -105.9414
12.9963 | -105.9376
13 | -105.9325
13.0037 | -105.9268
13.0074 | -105.9164
13.0111 | -105.9094
13.0148 | -105.9037
13.0185 | -105.8949
13.0222 | -105.8842
13.0259 | -105.8751
13.0296 | -105.8668
13.0333 | -105.8555
13.037 | -105.8472
13.0407 | -105.8409
13.0444 | -105.8325
13.0481 | -105.8225
13.0519 | -105.8121
13.0556 | -105.8015
13.0593 | -105.793
13.063 | -105.782
13.0667 | -105.7738
13.0704 | -105.7653
13.0741 | -105.7536
13.0778 | -105.7418
13.0815 | -105.7281
13.0852 | -105.7157
13.0889 | -105.7088
13.0926 | -105.7017
13.0963 | -105.6946
13.1 | -105.6858
13.1037 | -105.6806
13.1074 | -105.6777
13.1111 | -105.6734
13.1148 | -105.6708
13.1185 | -105.6742
13.1222 | -105.6763
13.1259 | -105.6741
13.1296 | -105.6743
13.1333 | -105.6778
13.137 | -105.6825
13.1407 | -105.6884
13.1444 | -105.6943
13.1481 | -105.6955
13.1519 | -105.7019
13.1556 | -105.707
13.1593 | -105.711
13.163 | -105.7164
13.1667 | -105.7259
13.1704 | -105.7358
13.1741 | -105.7448
13.1778 | -105.752
13.1815 | -105.7547
13.1852 | -105.7607
13.1889 | -105.7664
13.1926 | -105.7718
13.1963 | -105.7787
13.2 | -105.7806
13.2037 | -105.7829
13.2074 | -105.7875
13.2111 | -105.7904
13.2148 | -105.7915
13.2185 | -105.7933
13.2222 | -105.7928
13.2259 | -105.7879
13.2296 | -105.782
13.2333 | -105.778
13.237 | -105.7716
13.2407 | -105.7662
13.2444 | -105.7602
13.2481 | -105.7505
13.2519 | -105.7408
13.2556 | -105.7305
13.2593 | -105.7234
13.263 | -105.7198
13.2667 | -105.7139
13.2704 | -105.7078
13.2741 | -105.7003
13.2778 | -105.6915
13.2815 | -105.685
13.2852 | -105.6795
13.2889 | -105.6763
13.2926 | -105.6703
13.2963 | -105.6646
13.3 | -105.6624
13.3037 | -105.6605
13.3074 | -105.6527
13.3111 | -105.6492
13.3148 | -105.6524
13.3185 | -105.6533
13.3222 | -105.6537
13.3259 | -105.652
13.3296 | -105.6513
13.3333 | -105.6522
13.337 | -105.6557
13.3407 | -105.6572
13.3444 | -105.6627
13.3481 | -105.667
13.3519 | -105.6707
13.3556 | -105.6667
13.3593 | -105.6632
13.363 | -105.6617
13.3667 | -105.6594
13.3704 | -105.6587
13.3741 | -105.6591
13.3778 | -105.658
13.3815 | -105.6571
13.3852 | -105.6547
13.3889 | -105.6498
13.3926 | -105.6456
13.3963 | -105.64
13.4 | -105.6334
13.4037 | -105.6261
13.4074 | -105.6166
13.4111 | -105.6088
13.4148 | -105.604
13.4185 | -105.5994
13.4222 | -105.5938
13.4259 | -105.5879
13.4296 | -105.586
13.4333 | -105.5855
13.437 | -105.5861
13.4407 | -105.5896
13.4444 | -105.5942
13.4481 | -105.5996
13.4519 | -105.6057
13.4556 | -105.6132
13.4593 | -105.6229
13.463 | -105.6352
13.4667 | -105.6467
13.4704 | -105.6601
13.4741 | -105.6722
13.4778 | -105.6871
13.4815 | -105.7021
13.4852 | -105.7169
13.4889 | -105.731
13.4926 | -105.7465
13.4963 | -105.7625
13.5 | -105.7832
13.5037 | -105.8002
13.5074 | -105.8126
13.5111 | -105.829
13.5148 | -105.8446
13.5185 | -105.86
13.5222 | -105.8761
13.5259 | -105.8915
13.5296 | -105.9057
13.5333 | -105.9207
13.537 | -105.9336
13.5407 | -105.9511
13.5444 | -105.9684
13.5481 | -105.9873
13.5519 | -106.0102
13.5556 | -106.0289
13.5593 | -106.045
13.563 | -106.0584
13.5667 | -106.0726
13.5704 | -106.0901
13.5741 | -106.1014
13.5778 | -106.1101
13.5815 | -106.122
13.5852 | -106.1367
13.5889 | -106.1518
13.5926 | -106.1652
13.5963 | -106.1804
13.6 | -106.1928
13.6037 | -106.207
13.6074 | -106.2207
13.6111 | -106.2321
13.6148 | -106.2426
13.6185 | -106.2516
13.6222 | -106.2601
13.6259 | -106.2682
13.6296 | -106.2736
13.6333 | -106.2748
13.637 | -106.2717
13.6407 | -106.2709
13.6444 | -106.2702
13.6481 | -106.2678
13.6519 | -106.2639
13.6556 | -106.2589
13.6593 | -106.2532
13.663 | -106.2457
13.6667 | -106.2362
13.6704 | -106.2257
13.6741 | -106.2109
13.6778 | -106.1951
13.6815 | -106.1779
13.6852 | -106.1597
13.6889 | -106.1367
13.6926 | -106.1168
13.6963 | -106.1003
13.7 | -106.0849
13.7037 | -106.0722
13.7074 | -106.0563
13.7111 | -106.041
13.7148 | -106.0298
13.7185 | -106.018
13.7222 | -106.0052
13.7259 | -105.9955
13.7296 | -105.983
13.7333 | -105.9727
13.737 | -105.9614
13.7407 | -105.9508
13.7444 | -105.9406
13.7481 | -105.9253
13.7519 | -105.9131
13.7556 | -105.903
13.7593 | -105.8946
13.763 | -105.8832
13.7667 | -105.8666
13.7704 | -105.855
13.7741 | -105.8421
13.7778 | -105.8259
13.7815 | -105.8057
13.7852 | -105.7854
13.7889 | -105.7712
13.7926 | -105.7557
13.7963 | -105.7434
13.8 | -105.7329
13.8037 | -105.7221
13.8074 | -105.7101
13.8111 | -105.7043
13.8148 | -105.7028
13.8185 | -105.7036
13.8222 | -105.7058
13.8259 | -105.7099
13.8296 | -105.7201
13.8333 | -105.7353
13.837 | -105.7451
13.8407 | -105.7534
13.8444 | -105.7646
13.8481 | -105.7766
13.8519 | -105.7914
13.8556 | -105.8062
13.8593 | -105.8234
13.863 | -105.8392
13.8667 | -105.8538
13.8704 | -105.8668
13.8741 | -105.8771
13.8778 | -105.888
13.8815 | -105.8996
13.8852 | -105.9104
13.8889 | -105.906
13.8926 | -105.9085
13.8963 | -105.9073
13.9 | -105.9079
13.9037 | -105.906
13.9074 | -105.9058
13.9111 | -105.907
13.9148 | -105.9048
13.9185 | -105.9015
13.9222 | -105.8994
13.9259 | -105.8932
13.9296 | -105.8901
13.9333 | -105.8871
13.937 | -105.8835
13.9407 | -105.8798
13.9444 | -105.8754
13.9481 | -105.8728
13.9519 | -105.8701
13.9556 | -105.8647
13.9593 | -105.8512
13.963 | -105.8308
13.9667 | -105.8166
13.9704 | -105.8071
13.9741 | -105.795
13.9778 | -105.7807
13.9815 | -105.7679
13.9852 | -105.7556
13.9889 | -105.7385
13.9926 | -105.721
13.9963 | -105.7005
14 | -105.6757
14.0037 | -105.653
14.0074 | -105.6333
14.0111 | -105.61
14.0148 | -105.5837
14.0185 | -105.5567
14.0222 | -105.5306
14.0259 | -105.5044
14.0296 | -105.475
14.0333 | -105.4468
14.037 | -105.4269
14.0407 | -105.4098
14.0444 | -105.3922
14.0481 | -105.3771
14.0519 | -105.3613
14.0556 | -105.349
14.0593 | -105.3414
14.063 | -105.3382
14.0667 | -105.341
14.0704 | -105.3483
14.0741 | -105.3595
14.0778 | -105.3704
14.0815 | -105.3843
14.0852 | -105.4029
14.0889 | -105.4286
14.0926 | -105.457
14.0963 | -105.4861
14.1 | -105.5167
14.1037 | -105.5483
14.1074 | -105.5836
14.1111 | -105.621
14.1148 | -105.6556
14.1185 | -105.6914
14.1222 | -105.7298
14.1259 | -105.7654
14.1296 | -105.8003
14.1333 | -105.8312
14.137 | -105.8586
14.1407 | -105.8881
14.1444 | -105.9149
14.1481 | -105.9393
14.1519 | -105.9599
14.1556 | -105.9763
14.1593 | -105.9913
14.163 | -106.01
14.1667 | -106.0291
14.1704 | -106.0367
14.1741 | -106.0343
14.1778 | -106.0327
14.1815 | -106.0283
14.1852 | -106.0204
14.1889 | -106.0108
14.1926 | -105.9983
14.1963 | -105.9846
14.2 | -105.9742
14.2037 | -105.9647
14.2074 | -105.9515
14.2111 | -105.9381
14.2148 | -105.9267
14.2185 | -105.9162
14.2222 | -105.9057
14.2259 | -105.8983
14.2296 | -105.893
14.2333 | -105.8928
14.237 | -105.8946
14.2407 | -105.8936
14.2444 | -105.8936
14.2481 | -105.8965
14.2519 | -105.9037
14.2556 | -105.9154
14.2593 | -105.9278
14.263 | -105.9403
14.2667 | -105.952
14.2704 | -105.9626
14.2741 | -105.9773
14.2778 | -105.9928
14.2815 | -106.0112
14.2852 | -106.0287
14.2889 | -106.0437
14.2926 | -106.0552
14.2963 | -106.0681
14.3 | -106.0821
14.3037 | -106.094
14.3074 | -106.1027
14.3111 | -106.109
14.3148 | -106.1101
14.3185 | -106.1152
14.3222 | -106.1167
14.3259 | -106.1143
14.3296 | -106.1135
14.3333 | -106.1135
14.337 | -106.1097
14.3407 | -106.1042
14.3444 | -106.0935
14.3481 | -106.082
14.3519 | -106.0711
14.3556 | -106.0605
14.3593 | -106.05
14.363 | -106.0372
14.3667 | -106.0256
14.3704 | -106.0159
14.3741 | -106.0059
14.3778 | -105.9955
14.3815 | -105.9839
14.3852 | -105.9754
14.3889 | -105.9693
14.3926 | -105.9616
14.3963 | -105.9562
14.4 | -105.952
14.4037 | -105.95
14.4074 | -105.9482
14.4111 | -105.9485
14.4148 | -105.9501
14.4185 | -105.9505
14.4222 | -105.9533
14.4259 | -105.9602
14.4296 | -105.9666
14.4333 | -105.9717
14.437 | -105.9779
14.4407 | -105.9845
14.4444 | -105.9887
14.4481 | -105.9966
14.4519 | -106.0109
14.4556 | -106.0239
14.4593 | -106.0329
14.463 | -106.0423
14.4667 | -106.0508
14.4704 | -106.0573
14.4741 | -106.066
14.4778 | -106.0779
14.4815 | -106.0841
14.4852 | -106.0913
14.4889 | -106.098
14.4926 | -106.1056
14.4963 | -106.1117
14.5 | -106.1162
14.5037 | -106.1189
14.5074 | -106.1192
14.5111 | -106.1188
14.5148 | -106.117
14.5185 | -106.1152
14.5222 | -106.1119
14.5259 | -106.1006
14.5296 | -106.0906
14.5333 | -106.0799
14.537 | -106.0684
14.5407 | -106.0557
14.5444 | -106.0427
14.5481 | -106.0289
14.5519 | -106.0157
14.5556 | -106.003
14.5593 | -105.9874
14.563 | -105.971
14.5667 | -105.9581
14.5704 | -105.9485
14.5741 | -105.938
14.5778 | -105.9279
14.5815 | -105.9227
14.5852 | -105.9199
14.5889 | -105.9146
14.5926 | -105.9106
14.5963 | -105.9083
14.6 | -105.9084
14.6037 | -105.913
14.6074 | -105.9165
14.6111 | -105.9188
14.6148 | -105.9232
14.6185 | -105.9275
14.6222 | -105.9311
14.6259 | -105.9384
14.6296 | -105.9462
14.6333 | -105.9526
14.637 | -105.957
14.6407 | -105.9601
14.6444 | -105.962
14.6481 | -105.966
14.6519 | -105.9691
14.6556 | -105.973
14.6593 | -105.975
14.663 | -105.9742
14.6667 | -105.9752
14.6704 | -105.9789
14.6741 | -105.9787
14.6778 | -105.9805
14.6815 | -105.9855
14.6852 | -105.9887
14.6889 | -105.9885
14.6926 | -105.9942
14.6963 | -105.9987
14.7 | -106.0037
14.7037 | -106.0078
14.7074 | -106.0144
14.7111 | -106.0229
14.7148 | -106.0281
14.7185 | -106.0345
14.7222 | -106.0404
14.7259 | -106.0489
14.7296 | -106.0555
14.7333 | -106.063
14.737 | -106.0728
14.7407 | -106.0786
14.7444 | -106.083
14.7481 | -106.0855
14.7519 | -106.0864
14.7556 | -106.0884
14.7593 | -106.09
14.763 | -106.091
14.7667 | -106.093
14.7704 | -106.0946
14.7741 | -106.0971
14.7778 | -106.0949
14.7815 | -106.0879
14.7852 | -106.084
14.7889 | -106.0798
14.7926 | -106.0751
14.7963 | -106.0725
14.8 | -106.0725
14.8037 | -106.0697
14.8074 | -106.066
14.8111 | -106.0616
14.8148 | -106.0564
14.8185 | -106.055
14.8222 | -106.056
14.8259 | -106.0567
14.8296 | -106.0576
14.8333 | -106.058
14.837 | -106.0567
14.8407 | -106.0594
14.8444 | -106.0651
14.8481 | -106.0693
14.8519 | -106.0724
14.8556 | -106.0803
14.8593 | -106.0882
14.863 | -106.0964
14.8667 | -106.1047
14.8704 | -106.1127
14.8741 | -106.1209
14.8778 | -106.1318
14.8815 | -106.1447
14.8852 | -106.1618
14.8889 | -106.1761
14.8926 | -106.1856
14.8963 | -106.1975
14.9 | -106.2128
14.9037 | -106.2236
14.9074 | -106.2324
14.9111 | -106.2429
14.9148 | -106.2523
14.9185 | -106.2604
14.9222 | -106.2673
14.9259 | -106.2743
14.9296 | -106.2818
14.9333 | -106.289
14.937 | -106.2887
14.9407 | -106.2876
14.9444 | -106.289
14.9481 | -106.2863
14.9519 | -106.2828
14.9556 | -106.2807
14.9593 | -106.2721
14.963 | -106.2684
14.9667 | -106.2638
14.9704 | -106.2561
14.9741 | -106.247
14.9778 | -106.2376
14.9815 | -106.225
14.9852 | -106.2165
14.9889 | -106.2078
14.9926 | -106.1986
14.9963 | -106.1941
15 | -106.188
15.0037 | -106.1819
15.0074 | -106.1761
15.0111 | -106.1694
15.0148 | -106.1649
15.0185 | -106.1639
15.0222 | -106.1627
15.0259 | -106.1627
15.0296 | -106.1625
15.0333 | -106.1609
15.037 | -106.1586
15.0407 | -106.1587
15.0444 | -106.1601
15.0481 | -106.1622
15.0519 | -106.1647
15.0556 | -106.1642
15.0593 | -106.1662
15.063 | -106.1725
15.0667 | -106.179
15.0704 | -106.1829
15.0741 | -106.1864
15.0778 | -106.1911
15.0815 | -106.1957
15.0852 | -106.1973
15.0889 | -106.1971
15.0926 | -106.1961
15.0963 | -106.1941
15.1 | -106.1923
15.1037 | -106.1925
15.1074 | -106.195
15.1111 | -106.1929
15.1148 | -106.1894
15.1185 | -106.1871
15.1222 | -106.1861
15.1259 | -106.1808
15.1296 | -106.1732
15.1333 | -106.1659
15.137 | -106.1601
15.1407 | -106.1541
15.1444 | -106.1457
15.1481 | -106.1374
15.1519 | -106.1302
15.1556 | -106.1224
15.1593 | -106.1119
15.163 | -106.1038
15.1667 | -106.0994
15.1704 | -106.0898
15.1741 | -106.0789
15.1778 | -106.0698
15.1815 | -106.0636
15.1852 | -106.0587
15.1889 | -106.056
15.1926 | -106.0551
15.1963 | -106.0528
15.2 | -106.0492
15.2037 | -106.0473
15.2074 | -106.046
15.2111 | -106.044
15.2148 | -106.0406
15.2185 | -106.0401
15.2222 | -106.0402
15.2259 | -106.0385
15.2296 | -106.0418
15.2333 | -106.0448
15.237 | -106.0464
15.2407 | -106.049
15.2444 | -106.0503
15.2481 | -106.0511
15.2519 | -106.0541
15.2556 | -106.053
15.2593 | -106.0525
15.263 | -106.0521
15.2667 | -106.0532
15.2704 | -106.0544
15.2741 | -106.0537
15.2778 | -106.0488
15.2815 | -106.0425
15.2852 | -106.0395
15.2889 | -106.0339
15.2926 | -106.0256
15.2963 | -106.0179
15.3 | -106.0091
15.3037 | -106.003
15.3074 | -105.9986
15.3111 | -105.9942
15.3148 | -105.9898
15.3185 | -105.984
15.3222 | -105.9792
15.3259 | -105.9754
15.3296 | -105.975
15.3333 | -105.9735
15.337 | -105.9707
15.3407 | -105.9691
15.3444 | -105.9677
15.3481 | -105.9695
15.3519 | -105.9718
15.3556 | -105.9729
15.3593 | -105.9743
15.363 | -105.9788
15.3667 | -105.984
15.3704 | -105.9919
15.3741 | -105.9993
15.3778 | -106.0075
15.3815 | -106.0173
15.3852 | -106.0239
15.3889 | -106.0336
15.3926 | -106.0451
15.3963 | -106.0574
15.4 | -106.0696
15.4037 | -106.0805
15.4074 | -106.09
15.4111 | -106.0998
15.4148 | -106.1139
15.4185 | -106.1289
15.4222 | -106.1393
15.4259 | -106.1488
15.4296 | -106.1609
15.4333 | -106.1731
15.437 | -106.1835
15.4407 | -106.1942
15.4444 | -106.2067
15.4481 | -106.214
15.4519 | -106.2227
15.4556 | -106.2291
15.4593 | -106.235
15.463 | -106.2413
15.4667 | -106.2471
15.4704 | -106.2508
15.4741 | -106.2538
15.4778 | -106.2557
15.4815 | -106.2576
15.4852 | -106.2586
15.4889 | -106.2568
15.4926 | -106.2518
15.4963 | -106.2489
15.5 | -106.248
15.5037 | -106.2479
15.5074 | -106.2452
15.5111 | -106.2455
15.5148 | -106.2489
15.5185 | -106.2502
15.5222 | -106.251
15.5259 | -106.2481
15.5296 | -106.244
15.5333 | -106.2394
15.537 | -106.2302
15.5407 | -106.2231
15.5444 | -106.2169
15.5481 | -106.2088
15.5519 | -106.2001
15.5556 | -106.1928
15.5593 | -106.1867
15.563 | -106.1817
15.5667 | -106.1792
15.5704 | -106.1738
15.5741 | -106.1672
15.5778 | -106.1609
15.5815 | -106.1557
15.5852 | -106.1476
15.5889 | -106.1394
15.5926 | -106.1286
15.5963 | -106.1185
15.6 | -106.1083
15.6037 | -106.0962
15.6074 | -106.0874
15.6111 | -106.0789
15.6148 | -106.0713
15.6185 | -106.0601
15.6222 | -106.046
15.6259 | -106.0359
15.6296 | -106.028
15.6333 | -106.0192
15.637 | -106.0082
15.6407 | -105.9996
15.6444 | -105.9909
15.6481 | -105.9818
15.6519 | -105.9721
15.6556 | -105.9628
15.6593 | -105.953
15.663 | -105.9435
15.6667 | -105.933
15.6704 | -105.9236
15.6741 | -105.9161
15.6778 | -105.9112
15.6815 | -105.9061
15.6852 | -105.9018
15.6889 | -105.897
15.6926 | -105.8922
15.6963 | -105.8912
15.7 | -105.8891
15.7037 | -105.8869
15.7074 | -105.8844
15.7111 | -105.883
15.7148 | -105.8837
15.7185 | -105.8855
15.7222 | -105.8847
15.7259 | -105.8812
15.7296 | -105.876
15.7333 | -105.8718
15.737 | -105.8671
15.7407 | -105.8617
15.7444 | -105.8586
15.7481 | -105.8546
15.7519 | -105.8531
15.7556 | -105.8492
15.7593 | -105.8439
15.763 | -105.8431
15.7667 | -105.8376
15.7704 | -105.8305
15.7741 | -105.8251
15.7778 | -105.8192
15.7815 | -105.8143
15.7852 | -105.8109
15.7889 | -105.8061
15.7926 | -105.8006
15.7963 | -105.7957
15.8 | -105.7916
15.8037 | -105.7886
15.8074 | -105.7882
15.8111 | -105.7889
15.8148 | -105.7885
15.8185 | -105.7901
15.8222 | -105.7936
15.8259 | -105.798
15.8296 | -105.8037
15.8333 | -105.8074
15.837 | -105.8124
15.8407 | -105.82
15.8444 | -105.8297
15.8481 | -105.8395
15.8519 | -105.8454
15.8556 | -105.8537
15.8593 | -105.8649
15.863 | -105.8786
15.8667 | -105.8932
15.8704 | -105.91
15.8741 | -105.9288
15.8778 | -105.9468
15.8815 | -105.9637
15.8852 | -105.9804
15.8889 | -105.9995
15.8926 | -106.0211
15.8963 | -106.0383
15.9 | -106.0538
15.9037 | -106.0698
15.9074 | -106.0845
15.9111 | -106.0966
15.9148 | -106.1074
15.9185 | -106.1216
15.9222 | -106.1343
15.9259 | -106.1459
15.9296 | -106.1557
15.9333 | -106.165
15.937 | -106.174
15.9407 | -106.1827
15.9444 | -106.1881
15.9481 | -106.1935
15.9519 | -106.2001
15.9556 | -106.2058
15.9593 | -106.2086
15.963 | -106.2106
15.9667 | -106.2134
15.9704 | -106.2148
15.9741 | -106.2177
15.9778 | -106.2189
15.9815 | -106.221
15.9852 | -106.2218
15.9889 | -106.2214
15.9926 | -106.2247
15.9963 | -106.2265
16 | -106.227
16.0037 | -106.2287
16.0074 | -106.2276
16.0111 | -106.2254
16.0148 | -106.2246
16.0185 | -106.2254
16.0222 | -106.2254
16.0259 | -106.2243
16.0296 | -106.2248
16.0333 | -106.2271
16.037 | -106.2252
16.0407 | -106.2264
16.0444 | -106.2306
16.0481 | -106.2321
16.0519 | -106.2325
16.0556 | -106.233
16.0593 | -106.234
16.063 | -106.2368
16.0667 | -106.2352
16.0704 | -106.2325
16.0741 | -106.2308
16.0778 | -106.2319
16.0815 | -106.2341
16.0852 | -106.2306
16.0889 | -106.2226
16.0926 | -106.2153
16.0963 | -106.2089
16.1 | -106.2002
16.1037 | -106.1919
16.1074 | -106.184
16.1111 | -106.1763
16.1148 | -106.1675
16.1185 | -106.1542
16.1222 | -106.1403
16.1259 | -106.1305
16.1296 | -106.1214
16.1333 | -106.1118
16.137 | -106.101
16.1407 | -106.0903
16.1444 | -106.0812
16.1481 | -106.0677
16.1519 | -106.0524
16.1556 | -106.0378
16.1593 | -106.0258
16.163 | -106.0131
16.1667 | -106.0013
16.1704 | -105.9917
16.1741 | -105.9874
16.1778 | -105.9843
16.1815 | -105.9795
16.1852 | -105.9757
16.1889 | -105.9725
16.1926 | -105.97
16.1963 | -105.9702
16.2 | -105.9693
16.2037 | -105.9695
16.2074 | -105.9742
16.2111 | -105.9787
16.2148 | -105.9853
16.2185 | -105.9918
16.2222 | -105.9987
16.2259 | -106.0083
16.2296 | -106.0177
16.2333 | -106.026
16.237 | -106.0357
16.2407 | -106.0471
16.2444 | -106.06
16.2481 | -106.0738
16.2519 | -106.09
16.2556 | -106.1061
16.2593 | -106.121
16.263 | -106.1337
16.2667 | -106.1447
16.2704 | -106.1579
16.2741 | -106.1671
16.2778 | -106.1733
16.2815 | -106.1828
16.2852 | -106.1925
16.2889 | -106.2002
16.2926 | -106.2082
16.2963 | -106.2144
16.3 | -106.2225
16.3037 | -106.2269
16.3074 | -106.228
16.3111 | -106.2292
16.3148 | -106.2306
16.3185 | -106.2304
16.3222 | -106.2286
16.3259 | -106.2233
16.3296 | -106.2172
16.3333 | -106.2131
16.337 | -106.2073
16.3407 | -106.2012
16.3444 | -106.1952
16.3481 | -106.1903
16.3519 | -106.1865
16.3556 | -106.1816
16.3593 | -106.1735
16.363 | -106.165
16.3667 | -106.1583
16.3704 | -106.154
16.3741 | -106.1512
16.3778 | -106.1482
16.3815 | -106.1479
16.3852 | -106.1497
16.3889 | -106.1487
16.3926 | -106.1489
16.3963 | -106.1539
16.4 | -106.1612
16.4037 | -106.1651
16.4074 | -106.1714
16.4111 | -106.1791
16.4148 | -106.1907
16.4185 | -106.2017
16.4222 | -106.2126
16.4259 | -106.2257
16.4296 | -106.2381
16.4333 | -106.2558
16.437 | -106.274
16.4407 | -106.2871
16.4444 | -106.3005
16.4481 | -106.3153
16.4519 | -106.3307
16.4556 | -106.345
16.4593 | -106.3589
16.463 | -106.3725
16.4667 | -106.3832
16.4704 | -106.3931
16.4741 | -106.3999
16.4778 | -106.407
16.4815 | -106.4131
16.4852 | -106.4171
16.4889 | -106.4196
16.4926 | -106.4225
16.4963 | -106.4244
16.5 | -106.4261
16.5037 | -106.4252
16.5074 | -106.424
16.5111 | -106.4237
16.5148 | -106.4192
16.5185 | -106.4122
16.5222 | -106.4047
16.5259 | -106.3957
16.5296 | -106.3895
16.5333 | -106.3825
16.537 | -106.3774
16.5407 | -106.3745
16.5444 | -106.3711
16.5481 | -106.3676
16.5519 | -106.3654
16.5556 | -106.3634
16.5593 | -106.3614
16.563 | -106.3624
16.5667 | -106.3622
16.5704 | -106.3606
16.5741 | -106.3588
16.5778 | -106.3595
16.5815 | -106.359
16.5852 | -106.358
16.5889 | -106.3595
16.5926 | -106.361
16.5963 | -106.364
16.6 | -106.3685
16.6037 | -106.3746
16.6074 | -106.3824
16.6111 | -106.3895
16.6148 | -106.3977
16.6185 | -106.4049
16.6222 | -106.4129
16.6259 | -106.4237
16.6296 | -106.431
16.6333 | -106.436
16.637 | -106.4422
16.6407 | -106.4503
16.6444 | -106.4571
16.6481 | -106.4668
16.6519 | -106.4739
16.6556 | -106.479
16.6593 | -106.4857
16.663 | -106.4902
16.6667 | -106.4936
16.6704 | -106.4947
16.6741 | -106.4941
16.6778 | -106.4948
16.6815 | -106.4971
16.6852 | -106.4977
16.6889 | -106.4984
16.6926 | -106.4973
16.6963 | -106.4931
16.7 | -106.4892
16.7037 | -106.4878
16.7074 | -106.4851
16.7111 | -106.4812
16.7148 | -106.4729
16.7185 | -106.4625
16.7222 | -106.4552
16.7259 | -106.4508
16.7296 | -106.4447
16.7333 | -106.4403
16.737 | -106.4349
16.7407 | -106.432
16.7444 | -106.4297
16.7481 | -106.4269
16.7519 | -106.4254
16.7556 | -106.4229
16.7593 | -106.4217
16.763 | -106.4215
16.7667 | -106.4173
16.7704 | -106.4148
16.7741 | -106.4115
16.7778 | -106.4109
16.7815 | -106.4108
16.7852 | -106.4115
16.7889 | -106.4136
16.7926 | -106.4164
16.7963 | -106.4188
16.8 | -106.4214
16.8037 | -106.4228
16.8074 | -106.4204
16.8111 | -106.4198
16.8148 | -106.4204
16.8185 | -106.4203
16.8222 | -106.4211
16.8259 | -106.4234
16.8296 | -106.4237
16.8333 | -106.4242
16.837 | -106.4259
16.8407 | -106.425
16.8444 | -106.4231
16.8481 | -106.4226
16.8519 | -106.4251
16.8556 | -106.4247
16.8593 | -106.4245
16.863 | -106.4254
16.8667 | -106.4254
16.8704 | -106.4278
16.8741 | -106.4308
16.8778 | -106.4335
16.8815 | -106.4359
16.8852 | -106.4386
16.8889 | -106.4396
16.8926 | -106.4427
16.8963 | -106.4504
16.9 | -106.4579
16.9037 | -106.4649
16.9074 | -106.4712
16.9111 | -106.4765
16.9148 | -106.4831
16.9185 | -106.492
16.9222 | -106.4988
16.9259 | -106.5046
16.9296 | -106.5092
16.9333 | -106.5154
16.937 | -106.5215
16.9407 | -106.5264
16.9444 | -106.5342
16.9481 | -106.541
16.9519 | -106.5493
16.9556 | -106.5549
16.9593 | -106.5597
16.963 | -106.5635
16.9667 | -106.5661
16.9704 | -106.5691
16.9741 | -106.569
16.9778 | -106.5693
16.9815 | -106.5689
16.9852 | -106.5681
16.9889 | -106.5654
16.9926 | -106.5639
16.9963 | -106.5627
17 | -106.56
17.0037 | -106.5556
17.0074 | -106.5516
17.0111 | -106.5487
17.0148 | -106.5471
17.0185 | -106.5434
17.0222 | -106.5387
17.0259 | -106.5319
17.0296 | -106.5288
17.0333 | -106.5253
17.037 | -106.5191
17.0407 | -106.5138
17.0444 | -106.5111
17.0481 | -106.5134
17.0519 | -106.5131
17.0556 | -106.5087
17.0593 | -106.504
17.063 | -106.5023
17.0667 | -106.5011
17.0704 | -106.5005
17.0741 | -106.5004
17.0778 | -106.5
17.0815 | -106.5001
17.0852 | -106.4995
17.0889 | -106.4985
17.0926 | -106.4971
17.0963 | -106.5035
17.1 | -106.5074
17.1037 | -106.5114
17.1074 | -106.5149
17.1111 | -106.5204
17.1148 | -106.5247
17.1185 | -106.5291
17.1222 | -106.5314
17.1259 | -106.5354
17.1296 | -106.5402
17.1333 | -106.5457
17.137 | -106.5518
17.1407 | -106.5576
17.1444 | -106.563
17.1481 | -106.5697
17.1519 | -106.5766
17.1556 | -106.583
17.1593 | -106.5881
17.163 | -106.5904
17.1667 | -106.594
17.1704 | -106.5987
17.1741 | -106.6036
17.1778 | -106.6082
17.1815 | -106.6125
17.1852 | -106.6157
17.1889 | -106.6151
17.1926 | -106.6136
17.1963 | -106.6119
17.2 | -106.6123
17.2037 | -106.6108
17.2074 | -106.607
17.2111 | -106.6001
17.2148 | -106.592
17.2185 | -106.5819
17.2222 | -106.5724
17.2259 | -106.5636
17.2296 | -106.5526
17.2333 | -106.5419
17.237 | -106.5284
17.2407 | -106.5143
17.2444 | -106.5014
17.2481 | -106.4882
17.2519 | -106.4744
17.2556 | -106.4578
17.2593 | -106.4383
17.263 | -106.4184
17.2667 | -106.3988
17.2704 | -106.3828
17.2741 | -106.3653
17.2778 | -106.3503
17.2815 | -106.3377
17.2852 | -106.3244
17.2889 | -106.3094
17.2926 | -106.2923
17.2963 | -106.2789
17.3 | -106.2714
17.3037 | -106.2648
17.3074 | -106.2573
17.3111 | -106.2504
17.3148 | -106.2452
17.3185 | -106.2419
17.3222 | -106.2405
17.3259 | -106.2399
17.3296 | -106.2435
17.3333 | -106.2492
17.337 | -106.2528
17.3407 | -106.2594
17.3444 | -106.2677
17.3481 | -106.2787
17.3519 | -106.293
17.3556 | -106.308
17.3593 | -106.3241
17.363 | -106.3402
17.3667 | -106.357
17.3704 | -106.3773
17.3741 | -106.398
17.3778 | -106.4181
17.3815 | -106.438
17.3852 | -106.458
17.3889 | -106.4769
17.3926 | -106.4951
17.3963 | -106.5116
17.4 | -106.5313
17.4037 | -106.551
17.4074 | -106.5702
17.4111 | -106.5847
17.4148 | -106.596
17.4185 | -106.6063
17.4222 | -106.6152
17.4259 | -106.622
17.4296 | -106.6259
17.4333 | -106.6299
17.437 | -106.6336
17.4407 | -106.6364
17.4444 | -106.6361
17.4481 | -106.634
17.4519 | -106.6335
17.4556 | -106.6309
17.4593 | -106.6296
17.463 | -106.6251
17.4667 | -106.6203
17.4704 | -106.6101
17.4741 | -106.6006
17.4778 | -106.5929
17.4815 | -106.5823
17.4852 | -106.5744
17.4889 | -106.566
17.4926 | -106.5555
17.4963 | -106.5466
17.5 | -106.5377
17.5037 | -106.5272
17.5074 | -106.5167
17.5111 | -106.5046
17.5148 | -106.4927
17.5185 | -106.4787
17.5222 | -106.4649
17.5259 | -106.4518
17.5296 | -106.4401
17.5333 | -106.4279
17.537 | -106.4159
17.5407 | -106.4073
17.5444 | -106.3972
17.5481 | -106.3861
17.5519 | -106.3778
17.5556 | -106.3694
17.5593 | -106.3625
17.563 | -106.3583
17.5667 | -106.3571
17.5704 | -106.3564
17.5741 | -106.3573
17.5778 | -106.3556
17.5815 | -106.3557
17.5852 | -106.3619
17.5889 | -106.3639
17.5926 | -106.3674
17.5963 | -106.3708
17.6 | -106.3734
17.6037 | -106.3798
17.6074 | -106.3865
17.6111 | -106.3957
17.6148 | -106.4048
17.6185 | -106.4149
17.6222 | -106.4264
17.6259 | -106.4394
17.6296 | -106.4598
17.6333 | -106.4736
17.637 | -106.4891
17.6407 | -106.505
17.6444 | -106.5212
17.6481 | -106.5392
17.6519 | -106.5568
17.6556 | -106.5743
17.6593 | -106.5919
17.663 | -106.6144
17.6667 | -106.6362
17.6704 | -106.6548
17.6741 | -106.6734
17.6778 | -106.6939
17.6815 | -106.7145
17.6852 | -106.7332
17.6889 | -106.7503
17.6926 | -106.7662
17.6963 | -106.7832
17.7 | -106.7986
17.7037 | -106.8101
17.7074 | -106.8204
17.7111 | -106.8281
17.7148 | -106.8357
17.7185 | -106.8435
17.7222 | -106.8525
17.7259 | -106.8568
17.7296 | -106.8567
17.7333 | -106.8569
17.737 | -106.8532
17.7407 | -106.8496
17.7444 | -106.8443
17.7481 | -106.8377
17.7519 | -106.8319
17.7556 | -106.8249
17.7593 | -106.818
17.763 | -106.8098
17.7667 | -106.7984
17.7704 | -106.7885
17.7741 | -106.7756
17.7778 | -106.7617
17.7815 | -106.7485
17.7852 | -106.7367
17.7889 | -106.7247
17.7926 | -106.7098
17.7963 | -106.6979
17.8 | -106.6855
17.8037 | -106.6713
17.8074 | -106.6592
17.8111 | -106.6474
17.8148 | -106.6351
17.8185 | -106.6235
17.8222 | -106.6139
17.8259 | -106.6033
17.8296 | -106.592
17.8333 | -106.5834
17.837 | -106.5754
17.8407 | -106.5689
17.8444 | -106.5627
17.8481 | -106.5597
17.8519 | -106.5586
17.8556 | -106.5561
17.8593 | -106.5542
17.863 | -106.5551
17.8667 | -106.5565
17.8704 | -106.5596
17.8741 | -106.5629
17.8778 | -106.5674
17.8815 | -106.5746
17.8852 | -106.5818
17.8889 | -106.5907
17.8926 | -106.6006
17.8963 | -106.6129
17.9 | -106.627
17.9037 | -106.6426
17.9074 | -106.6559
17.9111 | -106.6735
17.9148 | -106.6922
17.9185 | -106.7104
17.9222 | -106.7306
17.9259 | -106.7511
17.9296 | -106.7702
17.9333 | -106.7881
17.937 | -106.8084
17.9407 | -106.8294
17.9444 | -106.8488
17.9481 | -106.8666
17.9519 | -106.8829
17.9556 | -106.9009
17.9593 | -106.92
17.963 | -106.9368
17.9667 | -106.9493
17.9704 | -106.9623
17.9741 | -106.9734
17.9778 | -106.9812
17.9815 | -106.9877
17.9852 | -106.9938
17.9889 | -106.9988
17.9926 | -107.0003
17.9963 | -106.9999
18 | -107.001
18.0037 | -107.0015
18.0074 | -106.9975
18.0111 | -106.9942
18.0148 | -106.9899
18.0185 | -106.9846
18.0222 | -106.9786
18.0259 | -106.9693
18.0296 | -106.9578
18.0333 | -106.9451
18.037 | -106.9333
18.0407 | -106.9219
18.0444 | -106.9117
18.0481 | -106.9008
18.0519 | -106.8902
18.0556 | -106.8782
18.0593 | -106.8672
18.063 | -106.8559
18.0667 | -106.8428
18.0704 | -106.8313
18.0741 | -106.8196
18.0778 | -106.8094
18.0815 | -106.8002
18.0852 | -106.7943
18.0889 | -106.7864
18.0926 | -106.7767
18.0963 | -106.7664
18.1 | -106.7573
18.1037 | -106.7479
18.1074 | -106.739
18.1111 | -106.7306
18.1148 | -106.7219
18.1185 | -106.7114
18.1222 | -106.7021
18.1259 | -106.6918
18.1296 | -106.6803
18.1333 | -106.669
18.137 | -106.6588
18.1407 | -106.6509
18.1444 | -106.6385
18.1481 | -106.6247
18.1519 | -106.6135
18.1556 | -106.6028
18.1593 | -106.5889
18.163 | -106.5789
18.1667 | -106.5653
18.1704 | -106.5516
18.1741 | -106.538
18.1778 | -106.5246
18.1815 | -106.5114
18.1852 | -106.4988
18.1889 | -106.4876
18.1926 | -106.4767
18.1963 | -106.4645
18.2 | -106.4509
18.2037 | -106.4414
18.2074 | -106.4334
18.2111 | -106.4247
18.2148 | -106.4139
18.2185 | -106.4048
18.2222 | -106.3975
18.2259 | -106.3943
18.2296 | -106.3905
18.2333 | -106.3868
18.237 | -106.3833
18.2407 | -106.3818
18.2444 | -106.3821
18.2481 | -106.3822
18.2519 | -106.3851
18.2556 | -106.39
18.2593 | -106.3922
18.263 | -106.395
18.2667 | -106.4008
18.2704 | -106.4095
18.2741 | -106.4189
18.2778 | -106.4294
18.2815 | -106.4417
18.2852 | -106.4553
18.2889 | -106.4708
18.2926 | -106.4859
18.2963 | -106.4986
18.3 | -106.5108
18.3037 | -106.5254
18.3074 | -106.5414
18.3111 | -106.5553
18.3148 | -106.5713
18.3185 | -106.5861
18.3222 | -106.6011
18.3259 | -106.6158
18.3296 | -106.6283
18.3333 | -106.6418
18.337 | -106.6526
18.3407 | -106.6632
18.3444 | -106.6735
18.3481 | -106.6827
18.3519 | -106.6916
18.3556 | -106.6999
18.3593 | -106.7082
18.363 | -106.7129
18.3667 | -106.7152
18.3704 | -106.7192
18.3741 | -106.7229
18.3778 | -106.728
18.3815 | -106.7298
18.3852 | -106.7326
18.3889 | -106.7375
18.3926 | -106.7385
18.3963 | -106.7386
18.4 | -106.7379
18.4037 | -106.74
18.4074 | -106.7428
18.4111 | -106.743
18.4148 | -106.7419
18.4185 | -106.7423
18.4222 | -106.7407
18.4259 | -106.7397
18.4296 | -106.7413
18.4333 | -106.7429
18.437 | -106.7448
18.4407 | -106.7459
18.4444 | -106.7459
18.4481 | -106.7475
18.4519 | -106.7499
18.4556 | -106.7524
18.4593 | -106.756
18.463 | -106.7608
18.4667 | -106.7629
18.4704 | -106.7641
18.4741 | -106.7645
18.4778 | -106.7684
18.4815 | -106.7728
18.4852 | -106.778
18.4889 | -106.782
18.4926 | -106.7846
18.4963 | -106.7862
18.5 | -106.7885
18.5037 | -106.7904
18.5074 | -106.7946
18.5111 | -106.7983
18.5148 | -106.8007
18.5185 | -106.8019
18.5222 | -106.8058
18.5259 | -106.8101
18.5296 | -106.814
18.5333 | -106.8156
18.537 | -106.8171
18.5407 | -106.8199
18.5444 | -106.8207
18.5481 | -106.8202
18.5519 | -106.8216
18.5556 | -106.8248
18.5593 | -106.828
18.563 | -106.8313
18.5667 | -106.832
18.5704 | -106.8322
18.5741 | -106.8322
18.5778 | -106.8317
18.5815 | -106.8321
18.5852 | -106.8329
18.5889 | -106.8337
18.5926 | -106.832
18.5963 | -106.8274
18.6 | -106.8252
18.6037 | -106.8259
18.6074 | -106.8242
18.6111 | -106.8233
18.6148 | -106.8242
18.6185 | -106.8251
18.6222 | -106.8236
18.6259 | -106.8247
18.6296 | -106.8268
18.6333 | -106.8274
18.637 | -106.8298
18.6407 | -106.8297
18.6444 | -106.8323
18.6481 | -106.8329
18.6519 | -106.832
18.6556 | -106.8355
18.6593 | -106.8361
18.663 | -106.8363
18.6667 | -106.8374
18.6704 | -106.8378
18.6741 | -106.8372
18.6778 | -106.8378
18.6815 | -106.8374
18.6852 | -106.8376
18.6889 | -106.8386
18.6926 | -106.8384
18.6963 | -106.8381
18.7 | -106.8375
18.7037 | -106.8356
18.7074 | -106.833
18.7111 | -106.8315
18.7148 | -106.8299
18.7185 | -106.8275
18.7222 | -106.8235
18.7259 | -106.8197
18.7296 | -106.8169
18.7333 | -106.8158
18.737 | -106.8167
18.7407 | -106.8189
18.7444 | -106.8173
18.7481 | -106.8145
18.7519 | -106.8124
18.7556 | -106.8123
18.7593 | -106.8143
18.763 | -106.8176
18.7667 | -106.8211
18.7704 | -106.8225
18.7741 | -106.8218
18.7778 | -106.8231
18.7815 | -106.8239
18.7852 | -106.8246
18.7889 | -106.8258
18.7926 | -106.8261
18.7963 | -106.8273
18.8 | -106.8284
18.8037 | -106.8288
18.8074 | -106.8301
18.8111 | -106.8329
18.8148 | -106.8366
18.8185 | -106.8369
18.8222 | -106.8364
18.8259 | -106.8351
18.8296 | -106.8345
18.8333 | -106.8335
18.837 | -106.8335
18.8407 | -106.8347
18.8444 | -106.8341
18.8481 | -106.8332
18.8519 | -106.8321
18.8556 | -106.8326
18.8593 | -106.8349
18.863 | -106.836
18.8667 | -106.836
18.8704 | -106.8346
18.8741 | -106.8327
18.8778 | -106.831
18.8815 | -106.83
18.8852 | -106.8304
18.8889 | -106.8305
18.8926 | -106.8296
18.8963 | -106.8292
18.9 | -106.8294
18.9037 | -106.8282
18.9074 | -106.8232
18.9111 | -106.8192
18.9148 | -106.8145
18.9185 | -106.8101
18.9222 | -106.8057
18.9259 | -106.8022
18.9296 | -106.799
18.9333 | -106.7952
18.937 | -106.7909
18.9407 | -106.7881
18.9444 | -106.7846
18.9481 | -106.7795
18.9519 | -106.7737
18.9556 | -106.769
18.9593 | -106.764
18.963 | -106.7587
18.9667 | -106.7547
18.9704 | -106.7504
18.9741 | -106.7469
18.9778 | -106.745
18.9815 | -106.7424
18.9852 | -106.7381
18.9889 | -106.7337
18.9926 | -106.7322
18.9963 | -106.7301
19 | -106.7273
19.0037 | -106.7246
19.0074 | -106.7197
19.0111 | -106.7141
19.0148 | -106.7115
19.0185 | -106.7055
19.0222 | -106.7003
19.0259 | -106.6955
19.0296 | -106.69
19.0333 | -106.6849
19.037 | -106.6808
19.0407 | -106.6788
19.0444 | -106.6764
19.0481 | -106.6705
19.0519 | -106.6669
19.0556 | -106.6626
19.0593 | -106.6559
19.063 | -106.6499
19.0667 | -106.6436
19.0704 | -106.6357
19.0741 | -106.628
19.0778 | -106.6194
19.0815 | -106.6121
19.0852 | -106.6058
19.0889 | -106.6015
19.0926 | -106.596
19.0963 | -106.5909
19.1 | -106.5852
19.1037 | -106.5775
19.1074 | -106.57
19.1111 | -106.5631
19.1148 | -106.5552
19.1185 | -106.5436
19.1222 | -106.5329
19.1259 | -106.5207
19.1296 | -106.5092
19.1333 | -106.5008
19.137 | -106.4917
19.1407 | -106.4839
19.1444 | -106.4781
19.1481 | -106.4707
19.1519 | -106.4615
19.1556 | -106.4518
19.1593 | -106.4458
19.163 | -106.4404
19.1667 | -106.4361
19.1704 | -106.4317
19.1741 | -106.4292
19.1778 | -106.4326
19.1815 | -106.4342
19.1852 | -106.4354
19.1889 | -106.4368
19.1926 | -106.4376
19.1963 | -106.4392
19.2 | -106.4404
19.2037 | -106.4423
19.2074 | -106.4434
19.2111 | -106.4429
19.2148 | -106.4441
19.2185 | -106.4488
19.2222 | -106.4547
19.2259 | -106.4582
19.2296 | -106.4624
19.2333 | -106.4696
19.237 | -106.4754
19.2407 | -106.4808
19.2444 | -106.4872
19.2481 | -106.4946
19.2519 | -106.5019
19.2556 | -106.5102
19.2593 | -106.5184
19.263 | -106.5271
19.2667 | -106.5346
19.2704 | -106.5421
19.2741 | -106.5514
19.2778 | -106.5604
19.2815 | -106.5729
19.2852 | -106.5852
19.2889 | -106.5942
19.2926 | -106.6058
19.2963 | -106.6172
19.3 | -106.6269
19.3037 | -106.6353
19.3074 | -106.6433
19.3111 | -106.6535
19.3148 | -106.6587
19.3185 | -106.6605
19.3222 | -106.6678
19.3259 | -106.6752
19.3296 | -106.6817
19.3333 | -106.6875
19.337 | -106.6926
19.3407 | -106.6973
19.3444 | -106.7024
19.3481 | -106.7042
19.3519 | -106.7049
19.3556 | -106.7053
19.3593 | -106.7055
19.363 | -106.7042
19.3667 | -106.7037
19.3704 | -106.7036
19.3741 | -106.7037
19.3778 | -106.7057
19.3815 | -106.7062
19.3852 | -106.7046
19.3889 | -106.7055
19.3926 | -106.7082
19.3963 | -106.708
19.4 | -106.706
19.4037 | -106.7035
19.4074 | -106.6985
19.4111 | -106.6946
19.4148 | -106.6902
19.4185 | -106.6863
19.4222 | -106.6822
19.4259 | -106.6775
19.4296 | -106.6754
19.4333 | -106.6716
19.437 | -106.6667
19.4407 | -106.6617
19.4444 | -106.657
19.4481 | -106.6495
19.4519 | -106.6407
19.4556 | -106.6339
19.4593 | -106.6286
19.463 | -106.6249
19.4667 | -106.6191
19.4704 | -106.6112
19.4741 | -106.6067
19.4778 | -106.5987
19.4815 | -106.5909
19.4852 | -106.5845
19.4889 | -106.5773
19.4926 | -106.5675
19.4963 | -106.5576
19.5 | -106.549
19.5037 | -106.5386
19.5074 | -106.5264
19.5111 | -106.5146
19.5148 | -106.5037
19.5185 | -106.493
19.5222 | -106.4807
19.5259 | -106.4681
19.5296 | -106.458
19.5333 | -106.4469
19.537 | -106.4351
19.5407 | -106.4218
19.5444 | -106.4096
19.5481 | -106.4004
19.5519 | -106.39
19.5556 | -106.3797
19.5593 | -106.369
19.563 | -106.3575
19.5667 | -106.3477
19.5704 | -106.3407
19.5741 | -106.3335
19.5778 | -106.3259
19.5815 | -106.319
19.5852 | -106.3123
19.5889 | -106.3049
19.5926 | -106.3016
19.5963 | -106.3018
19.6 | -106.2985
19.6037 | -106.2943
19.6074 | -106.2918
19.6111 | -106.2912
19.6148 | -106.2905
19.6185 | -106.2896
19.6222 | -106.2899
19.6259 | -106.2934
19.6296 | -106.2988
19.6333 | -106.3025
19.637 | -106.3054
19.6407 | -106.3106
19.6444 | -106.3146
19.6481 | -106.3209
19.6519 | -106.3273
19.6556 | -106.3334
19.6593 | -106.339
19.663 | -106.3454
19.6667 | -106.3511
19.6704 | -106.357
19.6741 | -106.3643
19.6778 | -106.3744
19.6815 | -106.3831
19.6852 | -106.3881
19.6889 | -106.3915
19.6926 | -106.3941
19.6963 | -106.3965
19.7 | -106.4008
19.7037 | -106.4058
19.7074 | -106.4048
19.7111 | -106.4019
19.7148 | -106.401
19.7185 | -106.397
19.7222 | -106.393
19.7259 | -106.3921
19.7296 | -106.3893
19.7333 | -106.3831
19.737 | -106.3749
19.7407 | -106.3648
19.7444 | -106.3563
19.7481 | -106.3476
19.7519 | -106.3342
19.7556 | -106.3233
19.7593 | -106.3129
19.763 | -106.3011
19.7667 | -106.2896
19.7704 | -106.2792
19.7741 | -106.2687
19.7778 | -106.256
19.7815 | -106.2424
19.7852 | -106.2318
19.7889 | -106.2227
19.7926 | -106.2117
19.7963 | -106.2014
19.8 | -106.1914
19.8037 | -106.1832
19.8074 | -106.1771
19.8111 | -106.169
19.8148 | -106.1614
19.8185 | -106.1566
19.8222 | -106.1535
19.8259 | -106.1503
19.8296 | -106.1456
19.8333 | -106.1407
19.837 | -106.1384
19.8407 | -106.1387
19.8444 | -106.1428
19.8481 | -106.1462
19.8519 | -106.1481
19.8556 | -106.1511
19.8593 | -106.1562
19.863 | -106.1643
19.8667 | -106.1723
19.8704 | -106.1796
19.8741 | -106.1878
19.8778 | -106.1973
19.8815 | -106.2057
19.8852 | -106.213
19.8889 | -106.2202
19.8926 | -106.2285
19.8963 | -106.2384
19.9 | -106.2481
19.9037 | -106.2576
19.9074 | -106.2665
19.9111 | -106.2753
19.9148 | -106.2853
19.9185 | -106.2959
19.9222 | -106.3033
19.9259 | -106.3116
19.9296 | -106.3177
19.9333 | -106.3239
19.937 | -106.3297
19.9407 | -106.3357
19.9444 | -106.3415
19.9481 | -106.3455
19.9519 | -106.3471
19.9556 | -106.3486
19.9593 | -106.3477
19.963 | -106.348
19.9667 | -106.3502
19.9704 | -106.3514
19.9741 | -106.3518
19.9778 | -106.35
19.9815 | -106.3452
19.9852 | -106.3415
19.9889 | -106.3408
19.9926 | -106.3397
19.9963 | -106.3379
20 | -106.3347
20.0037 | -106.3308
20.0074 | -106.3275
20.0111 | -106.3246
20.0148 | -106.3218
20.0185 | -106.3207
20.0222 | -106.319
20.0259 | -106.3162
20.0296 | -106.3152
20.0333 | -106.3172
20.037 | -106.3185
20.0407 | -106.3189
20.0444 | -106.3193
20.0481 | -106.3181
20.0519 | -106.3175
20.0556 | -106.3187
20.0593 | -106.3218
20.063 | -106.3248
20.0667 | -106.3292
20.0704 | -106.3318
20.0741 | -106.334
20.0778 | -106.3363
20.0815 | -106.3387
20.0852 | -106.341
20.0889 | -106.3445
20.0926 | -106.3479
20.0963 | -106.3499
20.1 | -106.3531
20.1037 | -106.3561
20.1074 | -106.3607
20.1111 | -106.3665
20.1148 | -106.3735
20.1185 | -106.3804
20.1222 | -106.3859
20.1259 | -106.39
20.1296 | -106.3932
20.1333 | -106.3968
20.137 | -106.4013
20.1407 | -106.4052
20.1444 | -106.4091
20.1481 | -106.4131
20.1519 | -106.4184
20.1556 | -106.4232
20.1593 | -106.4282
20.163 | -106.4325
20.1667 | -106.439
20.1704 | -106.4467
20.1741 | -106.4539
20.1778 | -106.4598
20.1815 | -106.4646
20.1852 | -106.4693
20.1889 | -106.4746
20.1926 | -106.4784
20.1963 | -106.4807
20.2 | -106.4838
20.2037 | -106.4866
20.2074 | -106.4913
20.2111 | -106.4948
20.2148 | -106.4985
20.2185 | -106.5029
20.2222 | -106.5066
20.2259 | -106.5089
20.2296 | -106.5092
20.2333 | -106.5091
20.237 | -106.5096
20.2407 | -106.5065
20.2444 | -106.5027
20.2481 | -106.4975
20.2519 | -106.492
20.2556 | -106.4869
20.2593 | -106.4792
20.263 | -106.4713
20.2667 | -106.4634
20.2704 | -106.4527
20.2741 | -106.4421
20.2778 | -106.431
20.2815 | -106.4173
20.2852 | -106.4063
20.2889 | -106.3945
20.2926 | -106.3791
20.2963 | -106.3633
20.3 | -106.3482
20.3037 | -106.3268
20.3074 | -106.3085
20.3111 | -106.2899
20.3148 | -106.2731
20.3185 | -106.2539
20.3222 | -106.2322
20.3259 | -106.211
20.3296 | -106.1916
20.3333 | -106.1734
20.337 | -106.1543
20.3407 | -106.1333
20.3444 | -106.1103
20.3481 | -106.0887
20.3519 | -106.0695
20.3556 | -106.0507
20.3593 | -106.0295
20.363 | -106.0079
20.3667 | -105.9874
20.3704 | -105.9655
20.3741 | -105.9453
20.3778 | -105.9247
20.3815 | -105.9062
20.3852 | -105.8914
20.3889 | -105.8779
20.3926 | -105.8645
20.3963 | -105.8495
20.4 | -105.8341
20.4037 | -105.8202
20.4074 | -105.8092
20.4111 | -105.7987
20.4148 | -105.7889
20.4185 | -105.7813
20.4222 | -105.7763
20.4259 | -105.7716
20.4296 | -105.7667
20.4333 | -105.7622
20.437 | -105.7598
20.4407 | -105.7576
20.4444 | -105.7557
20.4481 | -105.7551
20.4519 | -105.7556
20.4556 | -105.757
20.4593 | -105.7581
20.463 | -105.7594
20.4667 | -105.7625
20.4704 | -105.765
20.4741 | -105.7676
20.4778 | -105.7707
20.4815 | -105.7774
20.4852 | -105.7839
20.4889 | -105.788
20.4926 | -105.7926
20.4963 | -105.7977
20.5 | -105.8032
20.5037 | -105.8084
20.5074 | -105.8127
20.5111 | -105.8182
20.5148 | -105.825
20.5185 | -105.8327
20.5222 | -105.841
20.5259 | -105.8463
20.5296 | -105.8527
20.5333 | -105.8603
20.537 | -105.8657
20.5407 | -105.869
20.5444 | -105.8723
20.5481 | -105.8772
20.5519 | -105.8808
20.5556 | -105.8842
20.5593 | -105.8867
20.563 | -105.886
20.5667 | -105.8857
20.5704 | -105.8861
20.5741 | -105.8858
20.5778 | -105.8831
20.5815 | -105.8793
20.5852 | -105.8759
20.5889 | -105.8733
20.5926 | -105.8709
20.5963 | -105.8687
20.6 | -105.8667
20.6037 | -105.8643
20.6074 | -105.8611
20.6111 | -105.8574
20.6148 | -105.8536
20.6185 | -105.8487
20.6222 | -105.8424
20.6259 | -105.8365
20.6296 | -105.8317
20.6333 | -105.8259
20.637 | -105.8194
20.6407 | -105.8113
20.6444 | -105.8024
20.6481 | -105.7935
20.6519 | -105.7891
20.6556 | -105.7843
20.6593 | -105.7798
20.663 | -105.7757
20.6667 | -105.7687
20.6704 | -105.7609
20.6741 | -105.7508
20.6778 | -105.7398
20.6815 | -105.7302
20.6852 | -105.7221
20.6889 | -105.7147
20.6926 | -105.7081
20.6963 | -105.7002
20.7 | -105.6923
20.7037 | -105.6853
20.7074 | -105.6777
20.7111 | -105.6697
20.7148 | -105.6624
20.7185 | -105.6548
20.7222 | -105.648
20.7259 | -105.642
20.7296 | -105.633
20.7333 | -105.6255
20.737 | -105.6194
20.7407 | -105.6117
20.7444 | -105.6057
20.7481 | -105.5992
20.7519 | -105.5932
20.7556 | -105.5874
20.7593 | -105.5803
20.763 | -105.5748
20.7667 | -105.5701
20.7704 | -105.5659
20.7741 | -105.559
20.7778 | -105.5531
20.7815 | -105.548
20.7852 | -105.543
20.7889 | -105.5384
20.7926 | -105.5353
20.7963 | -105.5335
20.8 | -105.5321
20.8037 | -105.5299
20.8074 | -105.5251
20.8111 | -105.5207
20.8148 | -105.5187
20.8185 | -105.5179
20.8222 | -105.5167
20.8259 | -105.5154
20.8296 | -105.5157
20.8333 | -105.5168
20.837 | -105.516
20.8407 | -105.5166
20.8444 | -105.5163
20.8481 | -105.5169
20.8519 | -105.5178
20.8556 | -105.5184
20.8593 | -105.519
20.863 | -105.5188
20.8667 | -105.5186
20.8704 | -105.5191
20.8741 | -105.5209
20.8778 | -105.5244
20.8815 | -105.5273
20.8852 | -105.53
20.8889 | -105.5312
20.8926 | -105.5332
20.8963 | -105.5375
20.9 | -105.5416
20.9037 | -105.5469
20.9074 | -105.553
20.9111 | -105.5595
20.9148 | -105.5686
20.9185 | -105.5765
20.9222 | -105.583
20.9259 | -105.5907
20.9296 | -105.6011
20.9333 | -105.6143
20.937 | -105.6275
20.9407 | -105.639
20.9444 | -105.6499
20.9481 | -105.6599
20.9519 | -105.6718
20.9556 | -105.6822
20.9593 | -105.6932
20.963 | -105.7042
20.9667 | -105.7153
20.9704 | -105.7279
20.9741 | -105.7405
20.9778 | -105.7496
20.9815 | -105.7594
20.9852 | -105.7694
20.9889 | -105.7802
20.9926 | -105.7926
20.9963 | -105.8044
21 | -105.8147
21.0037 | -105.8252
21.0074 | -105.8357
21.0111 | -105.8457
21.0148 | -105.856
21.0185 | -105.8656
21.0222 | -105.8752
21.0259 | -105.8856
21.0296 | -105.8933
21.0333 | -105.9019
21.037 | -105.9093
21.0407 | -105.9128
21.0444 | -105.9179
21.0481 | -105.9251
21.0519 | -105.932
21.0556 | -105.9384
21.0593 | -105.9456
21.063 | -105.9518
21.0667 | -105.955
21.0704 | -105.9591
21.0741 | -105.9625
21.0778 | -105.9683
21.0815 | -105.9729
21.0852 | -105.9758
21.0889 | -105.976
21.0926 | -105.9732
21.0963 | -105.9704
21.1 | -105.9697
21.1037 | -105.971
21.1074 | -105.972
21.1111 | -105.9706
21.1148 | -105.9703
21.1185 | -105.968
21.1222 | -105.9627
21.1259 | -105.9591
21.1296 | -105.9565
21.1333 | -105.9539
21.137 | -105.9508
21.1407 | -105.946
21.1444 | -105.9423
21.1481 | -105.9391
21.1519 | -105.9343
21.1556 | -105.9302
21.1593 | -105.926
21.163 | -105.9219
21.1667 | -105.9179
21.1704 | -105.9116
21.1741 | -105.9063
21.1778 | -105.9009
21.1815 | -105.8948
21.1852 | -105.8866
21.1889 | -105.8823
21.1926 | -105.8784
21.1963 | -105.8722
21.2 | -105.8669
21.2037 | -105.8586
21.2074 | -105.8514
21.2111 | -105.8465
21.2148 | -105.841
21.2185 | -105.837
21.2222 | -105.8344
21.2259 | -105.8303
21.2296 | -105.826
21.2333 | -105.8209
21.237 | -105.8148
21.2407 | -105.8113
21.2444 | -105.8102
21.2481 | -105.809
21.2519 | -105.8068
21.2556 | -105.8064
21.2593 | -105.8037
21.263 | -105.8039
21.2667 | -105.8054
21.2704 | -105.8057
21.2741 | -105.8053
21.2778 | -105.807
21.2815 | -105.8062
21.2852 | -105.8046
21.2889 | -105.8012
21.2926 | -105.7976
21.2963 | -105.7956
21.3 | -105.7951
21.3037 | -105.7948
21.3074 | -105.792
21.3111 | -105.7904
21.3148 | -105.7888
21.3185 | -105.7886
21.3222 | -105.7871
21.3259 | -105.7856
21.3296 | -105.7821
21.3333 | -105.7763
21.337 | -105.7743
21.3407 | -105.7747
21.3444 | -105.774
21.3481 | -105.7701
21.3519 | -105.7652
21.3556 | -105.7627
21.3593 | -105.7578
21.363 | -105.7551
21.3667 | -105.753
21.3704 | -105.7499
21.3741 | -105.7473
21.3778 | -105.7446
21.3815 | -105.7407
21.3852 | -105.7368
21.3889 | -105.7322
21.3926 | -105.7304
21.3963 | -105.7285
21.4 | -105.7262
21.4037 | -105.7242
21.4074 | -105.7211
21.4111 | -105.7162
21.4148 | -105.7112
21.4185 | -105.707
21.4222 | -105.7039
21.4259 | -105.702
21.4296 | -105.6976
21.4333 | -105.6929
21.437 | -105.6898
21.4407 | -105.6871
21.4444 | -105.6844
21.4481 | -105.6808
21.4519 | -105.6775
21.4556 | -105.6745
21.4593 | -105.6735
21.463 | -105.6738
21.4667 | -105.6731
21.4704 | -105.6721
21.4741 | -105.6723
21.4778 | -105.6732
21.4815 | -105.6739
21.4852 | -105.6756
21.4889 | -105.6787
21.4926 | -105.6802
21.4963 | -105.6818
21.5 | -105.6844
21.5037 | -105.6873
21.5074 | -105.6906
21.5111 | -105.6936
21.5148 | -105.6955
21.5185 | -105.6983
21.5222 | -105.7004
21.5259 | -105.7011
21.5296 | -105.7024
21.5333 | -105.705
21.537 | -105.7063
21.5407 | -105.709
21.5444 | -105.7087
21.5481 | -105.7071
21.5519 | -105.707
21.5556 | -105.7064
21.5593 | -105.7056
21.563 | -105.7059
21.5667 | -105.706
21.5704 | -105.7046
21.5741 | -105.703
21.5778 | -105.704
21.5815 | -105.7031
21.5852 | -105.7009
21.5889 | -105.6989
21.5926 | -105.6977
21.5963 | -105.6937
21.6 | -105.6888
21.6037 | -105.6847
21.6074 | -105.6805
21.6111 | -105.6778
21.6148 | -105.6715
21.6185 | -105.6635
21.6222 | -105.6567
21.6259 | -105.6521
21.6296 | -105.6458
21.6333 | -105.6413
21.637 | -105.6367
21.6407 | -105.6271
21.6444 | -105.6173
21.6481 | -105.6075
21.6519 | -105.5989
21.6556 | -105.5914
21.6593 | -105.5821
21.663 | -105.5743
21.6667 | -105.565
21.6704 | -105.5569
21.6741 | -105.5493
21.6778 | -105.5418
21.6815 | -105.5338
21.6852 | -105.5246
21.6889 | -105.5161
21.6926 | -105.508
21.6963 | -105.5025
21.7 | -105.4939
21.7037 | -105.4864
21.7074 | -105.4796
21.7111 | -105.4703
21.7148 | -105.463
21.7185 | -105.4584
21.7222 | -105.4529
21.7259 | -105.4463
21.7296 | -105.4406
21.7333 | -105.4352
21.737 | -105.4324
21.7407 | -105.4297
21.7444 | -105.4265
21.7481 | -105.4227
21.7519 | -105.4186
21.7556 | -105.4166
21.7593 | -105.4146
21.763 | -105.4125
21.7667 | -105.41
21.7704 | -105.4083
21.7741 | -105.408
21.7778 | -105.4072
21.7815 | -105.406
21.7852 | -105.4031
21.7889 | -105.3993
21.7926 | -105.395
21.7963 | -105.3895
21.8 | -105.3818
21.8037 | -105.3767
21.8074 | -105.3737
21.8111 | -105.3681
21.8148 | -105.3642
21.8185 | -105.3607
21.8222 | -105.3564
21.8259 | -105.3492
21.8296 | -105.3426
21.8333 | -105.3364
21.837 | -105.3287
21.8407 | -105.3199
21.8444 | -105.3111
21.8481 | -105.3024
21.8519 | -105.2936
21.8556 | -105.2837
21.8593 | -105.2724
21.863 | -105.2611
21.8667 | -105.2492
21.8704 | -105.2383
21.8741 | -105.2249
21.8778 | -105.2117
21.8815 | -105.199
21.8852 | -105.1841
21.8889 | -105.1694
21.8926 | -105.1567
21.8963 | -105.1427
21.9 | -105.1282
21.9037 | -105.1152
21.9074 | -105.1012
21.9111 | -105.0875
21.9148 | -105.071
21.9185 | -105.056
21.9222 | -105.0418
21.9259 | -105.0297
21.9296 | -105.0172
21.9333 | -105.0024
21.937 | -104.9878
21.9407 | -104.9739
21.9444 | -104.9592
21.9481 | -104.9465
21.9519 | -104.9367
21.9556 | -104.9268
21.9593 | -104.9176
21.963 | -104.9086
21.9667 | -104.9006
21.9704 | -104.8928
21.9741 | -104.8847
21.9778 | -104.8791
21.9815 | -104.8752
21.9852 | -104.872
21.9889 | -104.8682
21.9926 | -104.8673
21.9963 | -104.8685
22 | -104.8685
22.0037 | -104.8693
22.0074 | -104.8702
22.0111 | -104.8739
22.0148 | -104.8773
22.0185 | -104.8817
22.0222 | -104.8876
22.0259 | -104.8948
22.0296 | -104.9007
22.0333 | -104.9077
22.037 | -104.9159
22.0407 | -104.9242
22.0444 | -104.9328
22.0481 | -104.9412
22.0519 | -104.9507
22.0556 | -104.9607
22.0593 | -104.9712
22.063 | -104.982
22.0667 | -104.9923
22.0704 | -105.003
22.0741 | -105.0129
22.0778 | -105.0236
22.0815 | -105.0339
22.0852 | -105.0446
22.0889 | -105.0555
22.0926 | -105.0674
22.0963 | -105.0809
22.1 | -105.0963
22.1037 | -105.1103
22.1074 | -105.1242
22.1111 | -105.1391
22.1148 | -105.1501
22.1185 | -105.161
22.1222 | -105.1752
22.1259 | -105.1853
22.1296 | -105.1942
22.1333 | -105.2043
22.137 | -105.2149
22.1407 | -105.225
22.1444 | -105.2339
22.1481 | -105.2436
22.1519 | -105.2521
22.1556 | -105.26
22.1593 | -105.2658
22.163 | -105.273
22.1667 | -105.2789
22.1704 | -105.2831
22.1741 | -105.289
22.1778 | -105.2925
22.1815 | -105.2931
22.1852 | -105.295
22.1889 | -105.2974
22.1926 | -105.2988
22.1963 | -105.2999
22.2 | -105.3029
22.2037 | -105.3048
22.2074 | -105.3071
22.2111 | -105.3086
22.2148 | -105.3077
22.2185 | -105.3051
22.2222 | -105.3046
22.2259 | -105.3054
22.2296 | -105.3077
22.2333 | -105.3078
22.237 | -105.3072
22.2407 | -105.306
22.2444 | -105.3054
22.2481 | -105.3067
22.2519 | -105.3077
22.2556 | -105.3089
22.2593 | -105.3118
22.263 | -105.3133
22.2667 | -105.3153
22.2704 | -105.3163
22.2741 | -105.3163
22.2778 | -105.3177
22.2815 | -105.3181
22.2852 | -105.3209
22.2889 | -105.3251
22.2926 | -105.3286
22.2963 | -105.3308
22.3 | -105.3338
22.3037 | -105.3383
22.3074 | -105.3411
22.3111 | -105.3434
22.3148 | -105.3487
22.3185 | -105.3535
22.3222 | -105.3569
22.3259 | -105.3607
22.3296 | -105.3658
22.3333 | -105.3681
22.337 | -105.3686
22.3407 | -105.3675
22.3444 | -105.3673
22.3481 | -105.3652
22.3519 | -105.3608
22.3556 | -105.3574
22.3593 | -105.3535
22.363 | -105.349
22.3667 | -105.3417
22.3704 | -105.3356
22.3741 | -105.3295
22.3778 | -105.3211
22.3815 | -105.3124
22.3852 | -105.3035
22.3889 | -105.2938
22.3926 | -105.2804
22.3963 | -105.2684
22.4 | -105.2545
22.4037 | -105.2408
22.4074 | -105.2242
22.4111 | -105.2052
22.4148 | -105.1843
22.4185 | -105.1638
22.4222 | -105.1394
22.4259 | -105.1134
22.4296 | -105.0866
22.4333 | -105.0566
22.437 | -105.0269
22.4407 | -104.987
22.4444 | -104.957
22.4481 | -104.9275
22.4519 | -104.8971
22.4556 | -104.8675
22.4593 | -104.837
22.463 | -104.8067
22.4667 | -104.7764
22.4704 | -104.7453
22.4741 | -104.7158
22.4778 | -104.6886
22.4815 | -104.6608
22.4852 | -104.6348
22.4889 | -104.6103
22.4926 | -104.5855
22.4963 | -104.5634
22.5 | -104.5436
22.5037 | -104.5242
22.5074 | -104.5053
22.5111 | -104.4891
22.5148 | -104.4743
22.5185 | -104.4593
22.5222 | -104.4477
22.5259 | -104.4385
22.5296 | -104.4319
22.5333 | -104.4287
22.537 | -104.4278
22.5407 | -104.4262
22.5444 | -104.4266
22.5481 | -104.4292
22.5519 | -104.4352
22.5556 | -104.4433
22.5593 | -104.4534
22.563 | -104.4651
22.5667 | -104.4784
22.5704 | -104.4929
22.5741 | -104.5098
22.5778 | -104.5287
22.5815 | -104.5477
22.5852 | -104.5672
22.5889 | -104.5889
22.5926 | -104.612
22.5963 | -104.6362
22.6 | -104.6623
22.6037 | -104.6891
22.6074 | -104.7158
22.6111 | -104.7432
22.6148 | -104.772
22.6185 | -104.8009
22.6222 | -104.8305
22.6259 | -104.8594
22.6296 | -104.8888
22.6333 | -104.9166
22.637 | -104.9456
22.6407 | -104.9756
22.6444 | -105.0028
22.6481 | -105.0287
22.6519 | -105.0563
22.6556 | -105.0848
22.6593 | -105.1128
22.663 | -105.1394
22.6667 | -105.1674
22.6704 | -105.1945
22.6741 | -105.2179
22.6778 | -105.2391
22.6815 | -105.2582
22.6852 | -105.2774
22.6889 | -105.294
22.6926 | -105.3095
22.6963 | -105.3247
22.7 | -105.3376
22.7037 | -105.3504
22.7074 | -105.3643
22.7111 | -105.377
22.7148 | -105.385
22.7185 | -105.3941
22.7222 | -105.4031
22.7259 | -105.413
22.7296 | -105.4235
22.7333 | -105.4337
22.737 | -105.4434
22.7407 | -105.4547
22.7444 | -105.4644
22.7481 | -105.474
22.7519 | -105.4851
22.7556 | -105.494
22.7593 | -105.5036
22.763 | -105.5118
22.7667 | -105.5171
22.7704 | -105.5212
22.7741 | -105.5261
22.7778 | -105.5307
22.7815 | -105.5367
22.7852 | -105.5428
22.7889 | -105.5458
22.7926 | -105.5509
22.7963 | -105.5541
22.8 | -105.5552
22.8037 | -105.5598
22.8074 | -105.564
22.8111 | -105.5685
22.8148 | -105.5725
22.8185 | -105.5764
22.8222 | -105.5793
22.8259 | -105.5825
22.8296 | -105.5854
22.8333 | -105.5886
22.837 | -105.5912
22.8407 | -105.5943
22.8444 | -105.5954
22.8481 | -105.5972
22.8519 | -105.5969
22.8556 | -105.5976
22.8593 | -105.5993
22.863 | -105.6009
22.8667 | -105.6006
22.8704 | -105.6003
22.8741 | -105.599
22.8778 | -105.5955
22.8815 | -105.5922
22.8852 | -105.5921
22.8889 | -105.5888
22.8926 | -105.5868
22.8963 | -105.5846
22.9 | -105.5792
22.9037 | -105.576
22.9074 | -105.5742
22.9111 | -105.5731
22.9148 | -105.5729
22.9185 | -105.5715
22.9222 | -105.5687
22.9259 | -105.5636
22.9296 | -105.5578
22.9333 | -105.5523
22.937 | -105.5473
22.9407 | -105.5438
22.9444 | -105.5403
22.9481 | -105.5384
22.9519 | -105.5354
22.9556 | -105.5315
22.9593 | -105.5262
22.963 | -105.5188
22.9667 | -105.514
22.9704 | -105.5095
22.9741 | -105.5034
22.9778 | -105.4995
22.9815 | -105.4946
22.9852 | -105.4889
22.9889 | -105.4837
22.9926 | -105.4765
22.9963 | -105.4709
23 | -105.4665
23.0037 | -105.4632
23.0074 | -105.4602
23.0111 | -105.4565
23.0148 | -105.4505
23.0185 | -105.4457
23.0222 | -105.4422
23.0259 | -105.4406
23.0296 | -105.4368
23.0333 | -105.4336
23.037 | -105.4321
23.0407 | -105.4317
23.0444 | -105.4301
23.0481 | -105.4288
23.0519 | -105.4272
23.0556 | -105.4257
23.0593 | -105.4255
23.063 | -105.4257
23.0667 | -105.4264
23.0704 | -105.4281
23.0741 | -105.4295
23.0778 | -105.4315
23.0815 | -105.4341
23.0852 | -105.4386
23.0889 | -105.4456
23.0926 | -105.4541
23.0963 | -105.4604
23.1 | -105.4647
23.1037 | -105.4683
23.1074 | -105.4715
23.1111 | -105.4765
23.1148 | -105.4831
23.1185 | -105.4887
23.1222 | -105.4944
23.1259 | -105.5012
23.1296 | -105.5054
23.1333 | -105.5105
23.137 | -105.5159
23.1407 | -105.5201
23.1444 | -105.5247
23.1481 | -105.5302
23.1519 | -105.5349
23.1556 | -105.539
23.1593 | -105.5402
23.163 | -105.5446
23.1667 | -105.5495
23.1704 | -105.552
23.1741 | -105.5528
23.1778 | -105.5567
23.1815 | -105.5579
23.1852 | -105.5587
23.1889 | -105.5609
23.1926 | -105.5632
23.1963 | -105.5649
23.2 | -105.5651
23.2037 | -105.5677
23.2074 | -105.5695
23.2111 | -105.5688
23.2148 | -105.5699
23.2185 | -105.5688
23.2222 | -105.5673
23.2259 | -105.5683
23.2296 | -105.5678
23.2333 | -105.5654
23.237 | -105.5613
23.2407 | -105.5561
23.2444 | -105.551
23.2481 | -105.5461
23.2519 | -105.5426
23.2556 | -105.5392
23.2593 | -105.5348
23.263 | -105.5295
23.2667 | -105.5225
23.2704 | -105.5152
23.2741 | -105.5082
23.2778 | -105.5006
23.2815 | -105.4913
23.2852 | -105.4827
23.2889 | -105.4764
23.2926 | -105.4701
23.2963 | -105.463
23.3 | -105.4543
23.3037 | -105.444
23.3074 | -105.434
23.3111 | -105.425
23.3148 | -105.4151
23.3185 | -105.4045
23.3222 | -105.3958
23.3259 | -105.3871
23.3296 | -105.3774
23.3333 | -105.367
23.337 | -105.3561
23.3407 | -105.3471
23.3444 | -105.3396
23.3481 | -105.3312
23.3519 | -105.3228
23.3556 | -105.3122
23.3593 | -105.3014
23.363 | -105.2911
23.3667 | -105.2799
23.3704 | -105.2705
23.3741 | -105.2631
23.3778 | -105.2554
23.3815 | -105.246
23.3852 | -105.2366
23.3889 | -105.2288
23.3926 | -105.223
23.3963 | -105.2182
23.4 | -105.2127
23.4037 | -105.2078
23.4074 | -105.2021
23.4111 | -105.197
23.4148 | -105.1917
23.4185 | -105.1878
23.4222 | -105.1846
23.4259 | -105.1839
23.4296 | -105.1855
23.4333 | -105.1875
23.437 | -105.1908
23.4407 | -105.1924
23.4444 | -105.194
23.4481 | -105.197
23.4519 | -105.2016
23.4556 | -105.2056
23.4593 | -105.2098
23.463 | -105.2129
23.4667 | -105.216
23.4704 | -105.2209
23.4741 | -105.2271
23.4778 | -105.233
23.4815 | -105.2408
23.4852 | -105.2494
23.4889 | -105.2592
23.4926 | -105.2663
23.4963 | -105.2739
23.5 | -105.2844
23.5037 | -105.2957
23.5074 | -105.3072
23.5111 | -105.3143
23.5148 | -105.323
23.5185 | -105.332
23.5222 | -105.3401
23.5259 | -105.3487
23.5296 | -105.357
23.5333 | -105.3657
23.537 | -105.3744
23.5407 | -105.3829
23.5444 | -105.3913
23.5481 | -105.3983
23.5519 | -105.4046
23.5556 | -105.4126
23.5593 | -105.4177
23.563 | -105.423
23.5667 | -105.4291
23.5704 | -105.4361
23.5741 | -105.4419
23.5778 | -105.4473
23.5815 | -105.4538
23.5852 | -105.4598
23.5889 | -105.4659
23.5926 | -105.4711
23.5963 | -105.4767
23.6 | -105.4826
23.6037 | -105.4881
23.6074 | -105.4922
23.6111 | -105.4984
23.6148 | -105.5069
23.6185 | -105.5119
23.6222 | -105.5147
23.6259 | -105.5154
23.6296 | -105.5166
23.6333 | -105.5222
23.637 | -105.5271
23.6407 | -105.5298
23.6444 | -105.5325
23.6481 | -105.5385
23.6519 | -105.5443
23.6556 | -105.5487
23.6593 | -105.5518
23.663 | -105.5561
23.6667 | -105.5599
23.6704 | -105.5638
23.6741 | -105.5684
23.6778 | -105.5733
23.6815 | -105.5758
23.6852 | -105.5789
23.6889 | -105.5835
23.6926 | -105.5895
23.6963 | -105.5943
23.7 | -105.6001
23.7037 | -105.6045
23.7074 | -105.6075
23.7111 | -105.6099
23.7148 | -105.6145
23.7185 | -105.6189
23.7222 | -105.6232
23.7259 | -105.6257
23.7296 | -105.6272
23.7333 | -105.6312
23.737 | -105.6353
23.7407 | -105.639
23.7444 | -105.6433
23.7481 | -105.6471
23.7519 | -105.6503
23.7556 | -105.6543
23.7593 | -105.6576
23.763 | -105.6597
23.7667 | -105.6618
23.7704 | -105.6641
23.7741 | -105.6688
23.7778 | -105.6743
23.7815 | -105.6791
23.7852 | -105.6821
23.7889 | -105.6872
23.7926 | -105.693
23.7963 | -105.7004
23.8 | -105.7066
23.8037 | -105.7124
23.8074 | -105.718
23.8111 | -105.7215
23.8148 | -105.724
23.8185 | -105.7279
23.8222 | -105.7322
23.8259 | -105.7378
23.8296 | -105.7431
23.8333 | -105.7465
23.837 | -105.751
23.8407 | -105.7563
23.8444 | -105.7624
23.8481 | -105.7694
23.8519 | -105.7741
23.8556 | -105.7788
23.8593 | -105.7842
23.863 | -105.79
23.8667 | -105.7952
23.8704 | -105.7997
23.8741 | -105.803
23.8778 | -105.8061
23.8815 | -105.8095
23.8852 | -105.8133
23.8889 | -105.8173
23.8926 | -105.8188
23.8963 | -105.8223
23.9 | -105.8244
23.9037 | -105.8252
23.9074 | -105.8264
23.9111 | -105.8276
23.9148 | -105.8298
23.9185 | -105.8309
23.9222 | -105.8301
23.9259 | -105.8293
23.9296 | -105.8291
23.9333 | -105.8267
23.937 | -105.8257
23.9407 | -105.8241
23.9444 | -105.8216
23.9481 | -105.8178
23.9519 | -105.8134
23.9556 | -105.8108
23.9593 | -105.8101
23.963 | -105.8112
23.9667 | -105.8107
23.9704 | -105.8059
23.9741 | -105.8023
23.9778 | -105.8001
23.9815 | -105.7941
23.9852 | -105.788
23.9889 | -105.7838
23.9926 | -105.7807
23.9963 | -105.7768
24 | -105.7724
24.0037 | -105.7689
24.0074 | -105.7651
24.0111 | -105.7599
24.0148 | -105.7536
24.0185 | -105.75
24.0222 | -105.7448
24.0259 | -105.7372
24.0296 | -105.7288
24.0333 | -105.7214
24.037 | -105.712
24.0407 | -105.7041
24.0444 | -105.6969
24.0481 | -105.6907
24.0519 | -105.6849
24.0556 | -105.6796
24.0593 | -105.6736
24.063 | -105.664
24.0667 | -105.6562
24.0704 | -105.6502
24.0741 | -105.6431
24.0778 | -105.6364
24.0815 | -105.6302
24.0852 | -105.6227
24.0889 | -105.6166
24.0926 | -105.6106
24.0963 | -105.6045
24.1 | -105.6
24.1037 | -105.5947
24.1074 | -105.5916
24.1111 | -105.5894
24.1148 | -105.5852
24.1185 | -105.581
24.1222 | -105.5776
24.1259 | -105.5769
24.1296 | -105.5755
24.1333 | -105.575
24.137 | -105.5749
24.1407 | -105.576
24.1444 | -105.5769
24.1481 | -105.5768
24.1519 | -105.5796
24.1556 | -105.5817
24.1593 | -105.5829
24.163 | -105.5832
24.1667 | -105.583
24.1704 | -105.5838
24.1741 | -105.5865
24.1778 | -105.5878
24.1815 | -105.5888
24.1852 | -105.591
24.1889 | -105.595
24.1926 | -105.5984
24.1963 | -105.6014
24.2 | -105.6055
24.2037 | -105.609
24.2074 | -105.6118
24.2111 | -105.615
24.2148 | -105.6187
24.2185 | -105.622
24.2222 | -105.6242
24.2259 | -105.6253
24.2296 | -105.6262
24.2333 | -105.6271
24.237 | -105.6282
24.2407 | -105.6291
24.2444 | -105.6278
24.2481 | -105.6186
24.2519 | -105.6095
24.2556 | -105.597
24.2593 | -105.5485
24.263 | -105.4423
24.2667 | -105.3071
24.2704 | -105.2073
24.2741 | -105.1047
24.2778 | -104.9852
24.2815 | -104.8595
24.2852 | -104.741
24.2889 | -104.6169
24.2926 | -104.5026
24.2963 | -104.3941
24.3 | -104.2784
24.3037 | -104.1563
24.3074 | -104.0301
24.3111 | -103.9012
24.3148 | -103.7684
24.3185 | -103.6315
24.3222 | -103.4917
24.3259 | -103.3475
24.3296 | -103.201
24.3333 | -103.0539
24.337 | -102.9065
24.3407 | -102.7588
24.3444 | -102.6103
24.3481 | -102.4622
24.3519 | -102.3161
24.3556 | -102.1721
24.3593 | -102.0295
24.363 | -101.8895
24.3667 | -101.7539
24.3704 | -101.6224
24.3741 | -101.4946
24.3778 | -101.3718
24.3815 | -101.2527
24.3852 | -101.1389
24.3889 | -101.0312
24.3926 | -100.9295
24.3963 | -100.8338
24.4 | -100.7443
24.4037 | -100.6621
24.4074 | -100.5874
24.4111 | -100.5188
24.4148 | -100.4572
24.4185 | -100.4031
24.4222 | -100.3562
24.4259 | -100.3167
24.4296 | -100.2855
24.4333 | -100.2621
24.437 | -100.2461
24.4407 | -100.2376
24.4444 | -100.2367
24.4481 | -100.2443
24.4519 | -100.2591
24.4556 | -100.2813
24.4593 | -100.3115
24.463 | -100.3488
24.4667 | -100.3932
24.4704 | -100.445
24.4741 | -100.5042
24.4778 | -100.5704
24.4815 | -100.6436
24.4852 | -100.724
24.4889 | -100.8111
24.4926 | -100.9044
24.4963 | -101.0038
24.5 | -101.1084
24.5037 | -101.2183
24.5074 | -101.3336
24.5111 | -101.4539
24.5148 | -101.5783
24.5185 | -101.7067
24.5222 | -101.8385
24.5259 | -101.9734
24.5296 | -102.1117
24.5333 | -102.2518
24.537 | -102.3942
24.5407 | -102.5378
24.5444 | -102.6809
24.5481 | -102.8235
24.5519 | -102.9668
24.5556 | -103.1079
24.5593 | -103.2425
24.563 | -103.3717
24.5667 | -103.4875
24.5704 | -103.5793
24.5741 | -103.6451
24.5778 | -103.5658
24.5815 | -103.5893
24.5852 | -103.5956
24.5889 | -103.6066
24.5926 | -103.6199
24.5963 | -103.6327
24.6 | -103.6348
24.6037 | -103.6239
24.6074 | -103.618
24.6111 | -103.6256
24.6148 | -103.6243
24.6185 | -103.5786
24.6222 | -103.5177
24.6259 | -103.4499
24.6296 | -103.3691
24.6333 | -103.2687
24.637 | -103.1584
24.6407 | -103.0306
24.6444 | -102.8913
24.6481 | -102.7512
24.6519 | -102.6101
24.6556 | -102.4679
24.6593 | -102.3246
24.663 | -102.1812
24.6667 | -102.0391
24.6704 | -101.9009
24.6741 | -101.7651
24.6778 | -101.6313
24.6815 | -101.5007
24.6852 | -101.3739
24.6889 | -101.2514
24.6926 | -101.1346
24.6963 | -101.0228
24.7 | -100.9157
24.7037 | -100.8141
24.7074 | -100.7182
24.7111 | -100.6286
24.7148 | -100.5447
24.7185 | -100.467
24.7222 | -100.3962
24.7259 | -100.3321
24.7296 | -100.2751
24.7333 | -100.2241
24.737 | -100.18
24.7407 | -100.143
24.7444 | -100.1134
24.7481 | -100.0914
24.7519 | -100.0762
24.7556 | -100.0681
24.7593 | -100.0679
24.763 | -100.0747
24.7667 | -100.0881
24.7704 | -100.1093
24.7741 | -100.1375
24.7778 | -100.1718
24.7815 | -100.2132
24.7852 | -100.2618
24.7889 | -100.3171
24.7926 | -100.3792
24.7963 | -100.4467
24.8 | -100.5203
24.8037 | -100.6001
24.8074 | -100.6851
24.8111 | -100.7762
24.8148 | -100.8724
24.8185 | -100.9722
24.8222 | -101.0759
24.8259 | -101.1839
24.8296 | -101.2961
24.8333 | -101.4118
24.837 | -101.5294
24.8407 | -101.6497
24.8444 | -101.772
24.8481 | -101.8974
24.8519 | -102.0246
24.8556 | -102.152
24.8593 | -102.2796
24.863 | -102.4069
24.8667 | -102.5332
24.8704 | -102.6592
24.8741 | -102.7843
24.8778 | -102.9066
24.8815 | -103.027
24.8852 | -103.1449
24.8889 | -103.2596
24.8926 | -103.372
24.8963 | -103.4811
24.9 | -103.585
24.9037 | -103.6865
24.9074 | -103.7879
24.9111 | -103.8871
24.9148 | -103.9907
24.9185 | -104.1055
24.9222 | -104.2292
24.9259 | -104.3484
24.9296 | -104.4521
24.9333 | -104.5589
24.937 | -104.6591
24.9407 | -104.7249
24.9444 | -104.7646
24.9481 | -104.7919
24.9519 | -104.8133
24.9556 | -104.8328
24.9593 | -104.8528
24.963 | -104.8746
24.9667 | -104.8972
24.9704 | -104.9143
24.9741 | -104.9266
24.9778 | -104.9383
24.9815 | -104.9489
24.9852 | -104.9588
24.9889 | -104.969
24.9926 | -104.9791
24.9963 | -104.9891
25 | -104.9977
25.0037 | -105.0066
25.0074 | -105.0151
25.0111 | -105.0239
25.0148 | -105.0349
25.0185 | -105.0469
25.0222 | -105.0569
25.0259 | -105.066
25.0296 | -105.074
25.0333 | -105.0798
25.037 | -105.085
25.0407 | -105.0902
25.0444 | -105.0935
25.0481 | -105.0964
25.0519 | -105.0983
25.0556 | -105.0987
25.0593 | -105.1002
25.063 | -105.1024
25.0667 | -105.1039
25.0704 | -105.1053
25.0741 | -105.1056
25.0778 | -105.1049
25.0815 | -105.1046
25.0852 | -105.1034
25.0889 | -105.1011
25.0926 | -105.0979
25.0963 | -105.0942
25.1 | -105.0905
25.1037 | -105.0868
25.1074 | -105.0831
25.1111 | -105.0771
25.1148 | -105.0735
25.1185 | -105.0698
25.1222 | -105.0673
25.1259 | -105.0641
25.1296 | -105.0604
25.1333 | -105.0576
25.137 | -105.0522
25.1407 | -105.0464
25.1444 | -105.0425
25.1481 | -105.0377
25.1519 | -105.034
25.1556 | -105.0313
25.1593 | -105.0266
25.163 | -105.0239
25.1667 | -105.0209
25.1704 | -105.0167
25.1741 | -105.0135
25.1778 | -105.0121
25.1815 | -105.011
25.1852 | -105.0078
25.1889 | -105.0044
25.1926 | -105.0046
25.1963 | -105.0032
25.2 | -105.0012
25.2037 | -104.9981
25.2074 | -104.9949
25.2111 | -104.9931
25.2148 | -104.9916
25.2185 | -104.9917
25.2222 | -104.9912
25.2259 | -104.99
25.2296 | -104.9865
25.2333 | -104.9857
25.237 | -104.985
25.2407 | -104.9832
25.2444 | -104.9815
25.2481 | -104.9801
25.2519 | -104.9789
25.2556 | -104.9784
25.2593 | -104.9762
25.263 | -104.9735
25.2667 | -104.9688
25.2704 | -104.9661
25.2741 | -104.9629
25.2778 | -104.9559
25.2815 | -104.9482
25.2852 | -104.9403
25.2889 | -104.9332
25.2926 | -104.9255
25.2963 | -104.9195
25.3 | -104.9119
25.3037 | -104.9007
25.3074 | -104.8912
25.3111 | -104.883
25.3148 | -104.8728
25.3185 | -104.8618
25.3222 | -104.8516
25.3259 | -104.8404
25.3296 | -104.8281
25.3333 | -104.8146
25.337 | -104.802
25.3407 | -104.7892
25.3444 | -104.777
25.3481 | -104.7647
25.3519 | -104.7538
25.3556 | -104.7439
25.3593 | -104.733
25.363 | -104.7207
25.3667 | -104.7082
25.3704 | -104.6941
25.3741 | -104.6792
25.3778 | -104.6645
25.3815 | -104.6522
25.3852 | -104.6411
25.3889 | -104.6316
25.3926 | -104.6212
25.3963 | -104.6093
25.4 | -104.5992
25.4037 | -104.5896
25.4074 | -104.5811
25.4111 | -104.5734
25.4148 | -104.5649
25.4185 | -104.5552
25.4222 | -104.5467
25.4259 | -104.5395
25.4296 | -104.5322
25.4333 | -104.5253
25.437 | -104.5192
25.4407 | -104.5136
25.4444 | -104.5075
25.4481 | -104.503
25.4519 | -104.4987
25.4556 | -104.4922
25.4593 | -104.4852
25.463 | -104.4795
25.4667 | -104.4756
25.4704 | -104.4718
25.4741 | -104.4688
25.4778 | -104.4671
25.4815 | -104.4643
25.4852 | -104.4606
25.4889 | -104.4582
25.4926 | -104.457
25.4963 | -104.4583
25.5 | -104.4602
25.5037 | -104.4599
25.5074 | -104.4596
25.5111 | -104.4602
25.5148 | -104.4599
25.5185 | -104.4622
25.5222 | -104.4651
25.5259 | -104.4684
25.5296 | -104.4723
25.5333 | -104.4772
25.537 | -104.4815
25.5407 | -104.4853
25.5444 | -104.4909
25.5481 | -104.4961
25.5519 | -104.5003
25.5556 | -104.5052
25.5593 | -104.5112
25.563 | -104.5183
25.5667 | -104.5253
25.5704 | -104.5322
25.5741 | -104.5395
25.5778 | -104.5456
25.5815 | -104.5525
25.5852 | -104.5595
25.5889 | -104.5683
25.5926 | -104.577
25.5963 | -104.5852
25.6 | -104.5935
25.6037 | -104.6031
25.6074 | -104.6128
25.6111 | -104.6232
25.6148 | -104.6328
25.6185 | -104.6441
25.6222 | -104.6549
25.6259 | -104.6655
25.6296 | -104.6753
25.6333 | -104.6867
25.637 | -104.7011
25.6407 | -104.7144
25.6444 | -104.7283
25.6481 | -104.7434
25.6519 | -104.7576
25.6556 | -104.7695
25.6593 | -104.7806
25.663 | -104.794
25.6667 | -104.8087
25.6704 | -104.8227
25.6741 | -104.8355
25.6778 | -104.8481
25.6815 | -104.8606
25.6852 | -104.8723
25.6889 | -104.8852
25.6926 | -104.8982
25.6963 | -104.9103
25.7 | -104.9226
25.7037 | -104.9338
25.7074 | -104.9446
25.7111 | -104.9561
25.7148 | -104.9661
25.7185 | -104.9746
25.7222 | -104.9837
25.7259 | -104.9958
25.7296 | -105.006
25.7333 | -105.0157
25.737 | -105.0242
25.7407 | -105.0317
25.7444 | -105.0391
25.7481 | -105.0461
25.7519 | -105.0506
25.7556 | -105.0552
25.7593 | -105.0605
25.763 | -105.0658
25.7667 | -105.0684
25.7704 | -105.0714
25.7741 | -105.0738
25.7778 | -105.0762
25.7815 | -105.0792
25.7852 | -105.0808
25.7889 | -105.0806
25.7926 | -105.0801
25.7963 | -105.0809
25.8 | -105.0815
25.8037 | -105.0825
25.8074 | -105.082
25.8111 | -105.0801
25.8148 | -105.0792
25.8185 | -105.0796
25.8222 | -105.0794
25.8259 | -105.0759
25.8296 | -105.0708
25.8333 | -105.0679
25.837 | -105.064
25.8407 | -105.0597
25.8444 | -105.0554
25.8481 | -105.0524
25.8519 | -105.0504
25.8556 | -105.0462
25.8593 | -105.0402
25.863 | -105.0346
25.8667 | -105.0283
25.8704 | -105.0223
25.8741 | -105.0176
25.8778 | -105.014
25.8815 | -105.0084
25.8852 | -105.0015
25.8889 | -104.997
25.8926 | -104.993
25.8963 | -104.9883
25.9 | -104.983
25.9037 | -104.9766
25.9074 | -104.9703
25.9111 | -104.9652
25.9148 | -104.9595
25.9185 | -104.9552
25.9222 | -104.9496
25.9259 | -104.9436
25.9296 | -104.9373
25.9333 | -104.9306
25.937 | -104.9234
25.9407 | -104.9169
25.9444 | -104.9113
25.9481 | -104.9058
25.9519 | -104.9003
25.9556 | -104.8929
25.9593 | -104.8853
25.963 | -104.8788
25.9667 | -104.8713
25.9704 | -104.8621
25.9741 | -104.8533
25.9778 | -104.8463
25.9815 | -104.8381
25.9852 | -104.8321
25.9889 | -104.8263
25.9926 | -104.8173
25.9963 | -104.8082
26 | -104.7982
26.0037 | -104.7904
26.0074 | -104.7813
26.0111 | -104.7724
26.0148 | -104.7616
26.0185 | -104.7512
26.0222 | -104.7418
26.0259 | -104.7329
26.0296 | -104.723
26.0333 | -104.7127
26.037 | -104.7034
26.0407 | -104.6962
26.0444 | -104.6885
26.0481 | -104.6796
26.0519 | -104.6688
26.0556 | -104.6576
26.0593 | -104.6506
26.063 | -104.6447
26.0667 | -104.6372
26.0704 | -104.6303
26.0741 | -104.6237
26.0778 | -104.6166
26.0815 | -104.6097
26.0852 | -104.6043
26.0889 | -104.5998
26.0926 | -104.5944
26.0963 | -104.5907
26.1 | -104.5877
26.1037 | -104.5855
26.1074 | -104.5834
26.1111 | -104.5814
26.1148 | -104.5802
26.1185 | -104.5794
26.1222 | -104.581
26.1259 | -104.5826
26.1296 | -104.5818
26.1333 | -104.5828
26.137 | -104.5827
26.1407 | -104.5844
26.1444 | -104.5872
26.1481 | -104.5907
26.1519 | -104.5945
26.1556 | -104.5986
26.1593 | -104.6036
26.163 | -104.6093
26.1667 | -104.6153
26.1704 | -104.6225
26.1741 | -104.6312
26.1778 | -104.6396
26.1815 | -104.6536
26.1852 | -104.6619
26.1889 | -104.6709
26.1926 | -104.6787
26.1963 | -104.6886
26.2 | -104.6989
26.2037 | -104.708
26.2074 | -104.7177
26.2111 | -104.7282
26.2148 | -104.7401
26.2185 | -104.7523
26.2222 | -104.7648
26.2259 | -104.7774
26.2296 | -104.7917
26.2333 | -104.8052
26.237 | -104.819
26.2407 | -104.8326
26.2444 | -104.8447
26.2481 | -104.8569
26.2519 | -104.8699
26.2556 | -104.8826
26.2593 | -104.8943
26.263 | -104.9061
26.2667 | -104.9178
26.2704 | -104.931
26.2741 | -104.9451
26.2778 | -104.9585
26.2815 | -104.9718
26.2852 | -104.9839
26.2889 | -104.994
26.2926 | -105.0045
26.2963 | -105.0137
26.3 | -105.0242
26.3037 | -105.0361
26.3074 | -105.0471
26.3111 | -105.0577
26.3148 | -105.0665
26.3185 | -105.0745
26.3222 | -105.0819
26.3259 | -105.089
26.3296 | -105.0965
26.3333 | -105.1029
26.337 | -105.1094
26.3407 | -105.1168
26.3444 | -105.1207
26.3481 | -105.1223
26.3519 | -105.1246
26.3556 | -105.1254
26.3593 | -105.1267
26.363 | -105.128
26.3667 | -105.1286
26.3704 | -105.1297
26.3741 | -105.1281
26.3778 | -105.1272
26.3815 | -105.1263
26.3852 | -105.1252
26.3889 | -105.1229
26.3926 | -105.1214
26.3963 | -105.118
26.4 | -105.1136
26.4037 | -105.1113
26.4074 | -105.1075
26.4111 | -105.1031
26.4148 | -105.1001
26.4185 | -105.0957
26.4222 | -105.0918
26.4259 | -105.0902
26.4296 | -105.0856
26.4333 | -105.08
26.437 | -105.0737
26.4407 | -105.0678
26.4444 | -105.0621
26.4481 | -105.0558
26.4519 | -105.0495
26.4556 | -105.0432
26.4593 | -105.0375
26.463 | -105.0329
26.4667 | -105.0289
26.4704 | -105.0233
26.4741 | -105.0189
26.4778 | -105.0155
26.4815 | -105.011
26.4852 | -105.0058
26.4889 | -105.0031
26.4926 | -105.001
26.4963 | -104.9981
26.5 | -104.9958
26.5037 | -104.9956
26.5074 | -104.9965
26.5111 | -104.9951
26.5148 | -104.9935
26.5185 | -104.9907
26.5222 | -104.9884
26.5259 | -104.9883
26.5296 | -104.9872
26.5333 | -104.9855
26.537 | -104.9854
26.5407 | -104.9861
26.5444 | -104.9857
26.5481 | -104.9835
26.5519 | -104.9811
26.5556 | -104.9803
26.5593 | -104.98
26.563 | -104.9792
26.5667 | -104.9814
26.5704 | -104.9825
26.5741 | -104.9824
26.5778 | -104.981
26.5815 | -104.9797
26.5852 | -104.9783
26.5889 | -104.9773
26.5926 | -104.9758
26.5963 | -104.9743
26.6 | -104.9742
26.6037 | -104.972
26.6074 | -104.9713
26.6111 | -104.972
26.6148 | -104.9724
26.6185 | -104.9719
26.6222 | -104.9717
26.6259 | -104.9726
26.6296 | -104.9728
26.6333 | -104.9702
26.637 | -104.9674
26.6407 | -104.9654
26.6444 | -104.9618
26.6481 | -104.957
26.6519 | -104.9529
26.6556 | -104.9478
26.6593 | -104.9458
26.663 | -104.944
26.6667 | -104.9405
26.6704 | -104.9363
26.6741 | -104.9294
26.6778 | -104.9221
26.6815 | -104.9163
26.6852 | -104.9096
26.6889 | -104.9041
26.6926 | -104.8993
26.6963 | -104.8936
26.7 | -104.8874
26.7037 | -104.8815
26.7074 | -104.8756
26.7111 | -104.8683
26.7148 | -104.8614
26.7185 | -104.8533
26.7222 | -104.8478
26.7259 | -104.843
26.7296 | -104.8349
26.7333 | -104.8264
26.737 | -104.8204
26.7407 | -104.814
26.7444 | -104.8053
26.7481 | -104.7955
26.7519 | -104.7882
26.7556 | -104.7808
26.7593 | -104.7702
26.763 | -104.7617
26.7667 | -104.7539
26.7704 | -104.7455
26.7741 | -104.7383
26.7778 | -104.7303
26.7815 | -104.7204
26.7852 | -104.7099
26.7889 | -104.6992
26.7926 | -104.6885
26.7963 | -104.6805
26.8 | -104.6726
26.8037 | -104.6613
26.8074 | -104.6507
26.8111 | -104.6405
26.8148 | -104.6318
26.8185 | -104.6231
26.8222 | -104.615
26.8259 | -104.6078
26.8296 | -104.6011
26.8333 | -104.594
26.837 | -104.5869
26.8407 | -104.5785
26.8444 | -104.5694
26.8481 | -104.5605
26.8519 | -104.5516
26.8556 | -104.5425
26.8593 | -104.5344
26.863 | -104.5288
26.8667 | -104.5219
26.8704 | -104.514
26.8741 | -104.5072
26.8778 | -104.501
26.8815 | -104.494
26.8852 | -104.4873
26.8889 | -104.4826
26.8926 | -104.4764
26.8963 | -104.4703
26.9 | -104.4651
26.9037 | -104.4595
26.9074 | -104.4558
26.9111 | -104.4522
26.9148 | -104.4476
26.9185 | -104.4433
26.9222 | -104.4407
26.9259 | -104.4389
26.9296 | -104.4362
26.9333 | -104.4329
26.937 | -104.4301
26.9407 | -104.4281
26.9444 | -104.4264
26.9481 | -104.4245
26.9519 | -104.4223
26.9556 | -104.4208
26.9593 | -104.42
26.963 | -104.4189
26.9667 | -104.4184
26.9704 | -104.4179
26.9741 | -104.4173
26.9778 | -104.416
26.9815 | -104.4164
26.9852 | -104.4153
26.9889 | -104.4143
26.9926 | -104.4131
26.9963 | -104.4144
27 | -104.4158
27.0037 | -104.4153
27.0074 | -104.4151
27.0111 | -104.4171
27.0148 | -104.4197
27.0185 | -104.4211
27.0222 | -104.4231
27.0259 | -104.4241
27.0296 | -104.425
27.0333 | -104.4274
27.037 | -104.4291
27.0407 | -104.4324
27.0444 | -104.4361
27.0481 | -104.4397
27.0519 | -104.4442
27.0556 | -104.4473
27.0593 | -104.4484
27.063 | -104.4503
27.0667 | -104.453
27.0704 | -104.4554
27.0741 | -104.4574
27.0778 | -104.4599
27.0815 | -104.4628
27.0852 | -104.4659
27.0889 | -104.4684
27.0926 | -104.4701
27.0963 | -104.4709
27.1 | -104.4737
27.1037 | -104.4746
27.1074 | -104.475
27.1111 | -104.4743
27.1148 | -104.474
27.1185 | -104.4716
27.1222 | -104.4683
27.1259 | -104.4677
27.1296 | -104.4666
27.1333 | -104.4639
27.137 | -104.4587
27.1407 | -104.453
27.1444 | -104.4486
27.1481 | -104.4428
27.1519 | -104.4371
27.1556 | -104.4306
27.1593 | -104.4257
27.163 | -104.4203
27.1667 | -104.4154
27.1704 | -104.41
27.1741 | -104.4035
27.1778 | -104.3953
27.1815 | -104.3867
27.1852 | -104.3798
27.1889 | -104.374
27.1926 | -104.3674
27.1963 | -104.3582
27.2 | -104.3468
27.2037 | -104.335
27.2074 | -104.3234
27.2111 | -104.3129
27.2148 | -104.3036
27.2185 | -104.2914
27.2222 | -104.2804
27.2259 | -104.271
27.2296 | -104.2616
27.2333 | -104.2528
27.237 | -104.2438
27.2407 | -104.2351
27.2444 | -104.225
27.2481 | -104.2183
27.2519 | -104.2099
27.2556 | -104.2006
27.2593 | -104.1922
27.263 | -104.1842
27.2667 | -104.1761
27.2704 | -104.1672
27.2741 | -104.159
27.2778 | -104.1522
27.2815 | -104.1462
27.2852 | -104.1399
27.2889 | -104.1323
27.2926 | -104.1242
27.2963 | -104.1172
27.3 | -104.1106
27.3037 | -104.1042
27.3074 | -104.0991
27.3111 | -104.0941
27.3148 | -104.0899
27.3185 | -104.0857
27.3222 | -104.081
27.3259 | -104.0784
27.3296 | -104.0753
27.3333 | -104.0721
27.337 | -104.0701
27.3407 | -104.0672
27.3444 | -104.0645
27.3481 | -104.0634
27.3519 | -104.0628
27.3556 | -104.062
27.3593 | -104.0598
27.363 | -104.0594
27.3667 | -104.06
27.3704 | -104.0611
27.3741 | -104.0631
27.3778 | -104.0653
27.3815 | -104.0664
27.3852 | -104.0675
27.3889 | -104.0707
27.3926 | -104.0727
27.3963 | -104.0741
27.4 | -104.0759
27.4037 | -104.0776
27.4074 | -104.0792
27.4111 | -104.0823
27.4148 | -104.0852
27.4185 | -104.0858
27.4222 | -104.0853
27.4259 | -104.0837
27.4296 | -104.0821
27.4333 | -104.0803
27.437 | -104.08
27.4407 | -104.0802
27.4444 | -104.079
27.4481 | -104.0781
27.4519 | -104.0761
27.4556 | -104.0732
27.4593 | -104.0715
27.463 | -104.0702
27.4667 | -104.0673
27.4704 | -104.0634
27.4741 | -104.0591
27.4778 | -104.0543
27.4815 | -104.0498
27.4852 | -104.0445
27.4889 | -104.0398
27.4926 | -104.0345
27.4963 | -104.0315
27.5 | -104.0279
27.5037 | -104.0224
27.5074 | -104.0183
27.5111 | -104.01
27.5148 | -104.0027
27.5185 | -103.9957
27.5222 | -103.9904
27.5259 | -103.9873
27.5296 | -103.981
27.5333 | -103.9732
27.537 | -103.9654
27.5407 | -103.959
27.5444 | -103.9535
27.5481 | -103.9468
27.5519 | -103.9404
27.5556 | -103.9342
27.5593 | -103.928
27.563 | -103.9222
27.5667 | -103.9169
27.5704 | -103.9113
27.5741 | -103.9037
27.5778 | -103.8982
27.5815 | -103.8945
27.5852 | -103.8923
27.5889 | -103.8901
27.5926 | -103.8887
27.5963 | -103.885
27.6 | -103.8805
27.6037 | -103.8794
27.6074 | -103.8783
27.6111 | -103.8763
27.6148 | -103.8755
27.6185 | -103.874
27.6222 | -103.8728
27.6259 | -103.8717
27.6296 | -103.8725
27.6333 | -103.8734
27.637 | -103.8741
27.6407 | -103.8749
27.6444 | -103.8767
27.6481 | -103.8791
27.6519 | -103.882
27.6556 | -103.8848
27.6593 | -103.8883
27.663 | -103.8911
27.6667 | -103.895
27.6704 | -103.898
27.6741 | -103.9004
27.6778 | -103.9047
27.6815 | -103.9092
27.6852 | -103.9128
27.6889 | -103.916
27.6926 | -103.9193
27.6963 | -103.923
27.7 | -103.9273
27.7037 | -103.931
27.7074 | -103.9335
27.7111 | -103.9363
27.7148 | -103.9382
27.7185 | -103.9403
27.7222 | -103.9433
27.7259 | -103.9471
27.7296 | -103.9505
27.7333 | -103.9512
27.737 | -103.9513
27.7407 | -103.9518
27.7444 | -103.9535
27.7481 | -103.9561
27.7519 | -103.9567
27.7556 | -103.957
27.7593 | -103.9562
27.763 | -103.9545
27.7667 | -103.9532
27.7704 | -103.9519
27.7741 | -103.9503
27.7778 | -103.9481
27.7815 | -103.941
27.7852 | -103.9379
27.7889 | -103.9348
27.7926 | -103.9298
27.7963 | -103.9265
27.8 | -103.9238
27.8037 | -103.9193
27.8074 | -103.9135
27.8111 | -103.9083
27.8148 | -103.9035
27.8185 | -103.899
27.8222 | -103.8934
27.8259 | -103.8877
27.8296 | -103.8796
27.8333 | -103.8717
27.837 | -103.8667
27.8407 | -103.8596
27.8444 | -103.8524
27.8481 | -103.8463
27.8519 | -103.8364
27.8556 | -103.8237
27.8593 | -103.8116
27.863 | -103.8017
27.8667 | -103.7933
27.8704 | -103.7859
27.8741 | -103.7771
27.8778 | -103.7684
27.8815 | -103.7585
27.8852 | -103.7499
27.8889 | -103.7393
27.8926 | -103.728
27.8963 | -103.7174
27.9 | -103.7073
27.9037 | -103.6993
27.9074 | -103.69
27.9111 | -103.6783
27.9148 | -103.6669
27.9185 | -103.6559
27.9222 | -103.6402
27.9259 | -103.6159
27.9296 | -103.5909
27.9333 | -103.5721
27.937 | -103.5553
27.9407 | -103.5399
27.9444 | -103.5236
27.9481 | -103.5085
27.9519 | -103.4938
27.9556 | -103.4794
27.9593 | -103.4646
27.963 | -103.4491
27.9667 | -103.434
27.9704 | -103.4174
27.9741 | -103.401
27.9778 | -103.3843
27.9815 | -103.3669
27.9852 | -103.3481
27.9889 | -103.3315
27.9926 | -103.313
27.9963 | -103.2924
28 | -103.2725
28.0037 | -103.2519