## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 286 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 291
Hood | 02-92B/P | 228
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1014.6 | hPa |  | PRC PRO 01
Humidity | 55.4 | %/Hr
Internal temperature of calibration bench | 24.09 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1020 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.3 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.03 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.02 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3067 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 11/09/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 13/09/2018
Technical Manager | Anthony HUE | Approval |  | 27/09/2018
CEO | Robert GLEMOT | Approval |  | 27/09/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3601 | 57.4243
0.012209 | -35.9412 | 38.226
0.018313 | -34.8853 | 28.0565
0.024417 | -34.4261 | 21.8036
0.030521 | -34.3131 | 18.0711
0.036626 | -34.1714 | 15.0192
0.04273 | -34.078 | 12.6573
0.048834 | -34.0359 | 11.1521
0.054938 | -33.9738 | 10.0206
0.061043 | -33.9156 | 8.93
0.067147 | -33.9054 | 8.0924
0.073251 | -33.9066 | 7.3215
0.079355 | -33.9248 | 6.6513
0.08546 | -33.9385 | 6.0549
0.091564 | -33.9347 | 5.6237
0.097668 | -33.9196 | 5.3823
0.10377 | -33.9015 | 5.1253
0.10988 | -33.8875 | 4.8455
0.11598 | -33.8804 | 4.5546
0.12209 | -33.8824 | 4.275
0.12819 | -33.891 | 3.9933
0.13429 | -33.8993 | 3.7082
0.1404 | -33.901 | 3.4407
0.1465 | -33.902 | 3.1913
0.15261 | -33.8995 | 2.9755
0.15871 | -33.8943 | 2.8115
0.16482 | -33.8877 | 2.7022
0.17092 | -33.8867 | 2.656
0.17702 | -33.8855 | 2.58
0.18313 | -33.8852 | 2.5007
0.18923 | -33.8848 | 2.3924
0.19534 | -33.8832 | 2.267
0.20144 | -33.8815 | 2.1255
0.20754 | -33.8788 | 2.0037
0.21365 | -33.8778 | 1.8894
0.21975 | -33.8775 | 1.7923
0.22586 | -33.8786 | 1.7075
0.23196 | -33.8809 | 1.6393
0.23807 | -33.8847 | 1.5803
0.24417 | -33.8888 | 1.522
0.25027 | -33.8929 | 1.4597
0.25638 | -33.8955 | 1.3969
0.26248 | -33.8977 | 1.328
0.26859 | -33.899 | 1.2593
0.27469 | -33.8997 | 1.1851
0.2808 | -33.9007 | 1.1113
0.2869 | -33.9017 | 1.0401
0.293 | -33.9026 | 0.97343
0.29911 | -33.9035 | 0.91401
0.30521 | -33.9045 | 0.85658
0.31132 | -33.9056 | 0.81924
0.31742 | -33.9073 | 0.77713
0.32353 | -33.9081 | 0.743
0.32963 | -33.9097 | 0.71446
0.33573 | -33.9115 | 0.69348
0.34184 | -33.9137 | 0.67101
0.34794 | -33.9167 | 0.64516
0.35405 | -33.9197 | 0.61991
0.36015 | -33.9229 | 0.59314
0.36626 | -33.9256 | 0.56367
0.37236 | -33.9283 | 0.52557
0.37846 | -33.931 | 0.47997
0.38457 | -33.9321 | 0.42243
0.39067 | -33.9334 | 0.37531
0.39678 | -33.9342 | 0.32333
0.40288 | -33.9347 | 0.26954
0.40899 | -33.9347 | 0.21783
0.41509 | -33.9341 | 0.16818
0.42119 | -33.9341 | 0.12708
0.4273 | -33.9341 | 0.082338
0.4334 | -33.934 | 0.03786
0.43951 | -33.9344 | -0.0034911
0.44561 | -33.9353 | -0.041515
0.45172 | -33.937 | -0.071366
0.45782 | -33.9379 | -0.10636
0.46392 | -33.9393 | -0.141
0.47003 | -33.9405 | -0.17146
0.47613 | -33.9415 | -0.20448
0.48224 | -33.9422 | -0.2382
0.48834 | -33.9416 | -0.27362
0.49445 | -33.9415 | -0.30814
0.50055 | -33.9412 | -0.34048
0.50665 | -33.9406 | -0.37501
0.51276 | -33.9398 | -0.40987
0.51886 | -33.9389 | -0.44477
0.52497 | -33.9378 | -0.47029
0.53107 | -33.9371 | -0.50099
0.53717 | -33.9365 | -0.53009
0.54328 | -33.9357 | -0.55447
0.54938 | -33.9355 | -0.57804
0.55549 | -33.9356 | -0.59987
0.56159 | -33.9366 | -0.61212
0.5677 | -33.9369 | -0.62751
0.5738 | -33.9379 | -0.64463
0.5799 | -33.9386 | -0.66011
0.58601 | -33.9396 | -0.67049
0.59211 | -33.9411 | -0.67672
0.59822 | -33.9422 | -0.68639
0.60432 | -33.9436 | -0.69906
0.61043 | -33.9447 | -0.707
0.61653 | -33.9458 | -0.71742
0.62263 | -33.9472 | -0.72639
0.62874 | -33.9478 | -0.73683
0.63484 | -33.9486 | -0.74574
0.64095 | -33.949 | -0.75571
0.64705 | -33.9494 | -0.76542
0.65316 | -33.9497 | -0.77766
0.65926 | -33.9499 | -0.79305
0.66536 | -33.9491 | -0.8105
0.67147 | -33.9489 | -0.82386
0.67757 | -33.9488 | -0.83833
0.68368 | -33.9483 | -0.8574
0.68978 | -33.9479 | -0.87464
0.69589 | -33.9473 | -0.89395
0.70199 | -33.9467 | -0.91701
0.70809 | -33.9462 | -0.93605
0.7142 | -33.9458 | -0.95676
0.7203 | -33.9451 | -0.97634
0.72641 | -33.9448 | -0.99908
0.73251 | -33.9445 | -1.0193
0.73862 | -33.9449 | -1.0413
0.74472 | -33.9448 | -1.0614
0.75082 | -33.9448 | -1.0838
0.75693 | -33.9454 | -1.1057
0.76303 | -33.9461 | -1.1278
0.76914 | -33.948 | -1.1481
0.77524 | -33.9491 | -1.1675
0.78135 | -33.9505 | -1.1876
0.78745 | -33.9519 | -1.2074
0.79355 | -33.9532 | -1.2273
0.79966 | -33.955 | -1.2466
0.80576 | -33.9572 | -1.2661
0.81187 | -33.9588 | -1.2874
0.81797 | -33.9608 | -1.3086
0.82408 | -33.9625 | -1.3291
0.83018 | -33.9642 | -1.3494
0.83628 | -33.9662 | -1.3678
0.84239 | -33.9674 | -1.3883
0.84849 | -33.9688 | -1.41
0.8546 | -33.9703 | -1.4295
0.8607 | -33.9715 | -1.4503
0.86681 | -33.9726 | -1.469
0.87291 | -33.9737 | -1.4878
0.87901 | -33.9742 | -1.506
0.88512 | -33.975 | -1.5241
0.89122 | -33.9758 | -1.5423
0.89733 | -33.9763 | -1.5581
0.90343 | -33.9768 | -1.5755
0.90953 | -33.977 | -1.5898
0.91564 | -33.9771 | -1.6055
0.92174 | -33.9774 | -1.6205
0.92785 | -33.9776 | -1.635
0.93395 | -33.9778 | -1.6485
0.94006 | -33.978 | -1.6614
0.94616 | -33.9779 | -1.6731
0.95226 | -33.9781 | -1.6857
0.95837 | -33.9783 | -1.6993
0.96447 | -33.9786 | -1.7118
0.97058 | -33.9785 | -1.7249
0.97668 | -33.9784 | -1.7379
0.98279 | -33.9783 | -1.7504
0.98889 | -33.9782 | -1.7609
0.99499 | -33.9783 | -1.774
1.0011 | -33.9783 | -1.7845
1.0072 | -33.9781 | -1.796
1.0133 | -33.9779 | -1.8065
1.0194 | -33.9777 | -1.8193
1.0255 | -33.9775 | -1.8328
1.0316 | -33.9771 | -1.8446
1.0377 | -33.9769 | -1.8581
1.0438 | -33.9765 | -1.8706
1.0499 | -33.976 | -1.8855
1.056 | -33.9754 | -1.8998
1.0621 | -33.975 | -1.914
1.0682 | -33.9747 | -1.9263
1.0743 | -33.9743 | -1.9416
1.0805 | -33.9739 | -1.9558
1.0866 | -33.9732 | -1.971
1.0927 | -33.9725 | -1.9855
1.0988 | -33.9721 | -2.0002
1.1049 | -33.9715 | -2.0143
1.111 | -33.9708 | -2.0296
1.1171 | -33.9703 | -2.0442
1.1232 | -33.97 | -2.0594
1.1293 | -33.9697 | -2.0735
1.1354 | -33.9696 | -2.0883
1.1415 | -33.9694 | -2.1039
1.1476 | -33.9689 | -2.1195
1.1537 | -33.9685 | -2.1344
1.1598 | -33.9687 | -2.1493
1.1659 | -33.9686 | -2.1628
1.172 | -33.9684 | -2.1768
1.1781 | -33.9683 | -2.1908
1.1842 | -33.9682 | -2.205
1.1903 | -33.9685 | -2.216
1.1964 | -33.9685 | -2.2296
1.2025 | -33.9687 | -2.2436
1.2086 | -33.969 | -2.2562
1.2147 | -33.9691 | -2.2688
1.2209 | -33.9695 | -2.28
1.227 | -33.9701 | -2.2907
1.2331 | -33.9705 | -2.3029
1.2392 | -33.9709 | -2.3147
1.2453 | -33.9713 | -2.3258
1.2514 | -33.9717 | -2.3357
1.2575 | -33.9722 | -2.3465
1.2636 | -33.973 | -2.3553
1.2697 | -33.9736 | -2.3645
1.2758 | -33.9743 | -2.3756
1.2819 | -33.9749 | -2.3854
1.288 | -33.9756 | -2.3961
1.2941 | -33.9764 | -2.4046
1.3002 | -33.9771 | -2.4127
1.3063 | -33.9776 | -2.4215
1.3124 | -33.9782 | -2.4305
1.3185 | -33.9789 | -2.4394
1.3246 | -33.9796 | -2.4482
1.3307 | -33.9804 | -2.4567
1.3368 | -33.9812 | -2.4653
1.3429 | -33.9819 | -2.4741
1.349 | -33.9824 | -2.4834
1.3551 | -33.9831 | -2.4925
1.3613 | -33.9837 | -2.5007
1.3674 | -33.9843 | -2.5106
1.3735 | -33.9848 | -2.5203
1.3796 | -33.9852 | -2.5293
1.3857 | -33.9856 | -2.5387
1.3918 | -33.9862 | -2.5487
1.3979 | -33.9868 | -2.5578
1.404 | -33.9873 | -2.5693
1.4101 | -33.9877 | -2.5796
1.4162 | -33.9882 | -2.5903
1.4223 | -33.9887 | -2.6016
1.4284 | -33.9892 | -2.6117
1.4345 | -33.9895 | -2.6232
1.4406 | -33.9899 | -2.6368
1.4467 | -33.9904 | -2.6484
1.4528 | -33.9909 | -2.6599
1.4589 | -33.9914 | -2.6718
1.465 | -33.9917 | -2.6827
1.4711 | -33.9922 | -2.6942
1.4772 | -33.9928 | -2.7076
1.4833 | -33.9933 | -2.7208
1.4894 | -33.9939 | -2.733
1.4955 | -33.9945 | -2.745
1.5016 | -33.9951 | -2.7577
1.5078 | -33.9956 | -2.7718
1.5139 | -33.996 | -2.7849
1.52 | -33.9965 | -2.7976
1.5261 | -33.9969 | -2.8104
1.5322 | -33.9974 | -2.8244
1.5383 | -33.9978 | -2.8371
1.5444 | -33.9984 | -2.85
1.5505 | -33.999 | -2.8636
1.5566 | -33.9996 | -2.8767
1.5627 | -34.0001 | -2.8906
1.5688 | -34.0008 | -2.9043
1.5749 | -34.0012 | -2.9173
1.581 | -34.0016 | -2.9308
1.5871 | -34.0021 | -2.9447
1.5932 | -34.0025 | -2.9578
1.5993 | -34.003 | -2.9721
1.6054 | -34.0035 | -2.9855
1.6115 | -34.004 | -2.9997
1.6176 | -34.0044 | -3.0124
1.6237 | -34.0047 | -3.0262
1.6298 | -34.005 | -3.0389
1.6359 | -34.0055 | -3.0518
1.642 | -34.006 | -3.0638
1.6482 | -34.0063 | -3.0769
1.6543 | -34.0068 | -3.0905
1.6604 | -34.0072 | -3.1027
1.6665 | -34.0074 | -3.1153
1.6726 | -34.0077 | -3.1279
1.6787 | -34.0081 | -3.1406
1.6848 | -34.0081 | -3.1533
1.6909 | -34.0084 | -3.1661
1.697 | -34.0088 | -3.1785
1.7031 | -34.0089 | -3.1912
1.7092 | -34.009 | -3.2034
1.7153 | -34.0092 | -3.2154
1.7214 | -34.0091 | -3.2276
1.7275 | -34.0091 | -3.2402
1.7336 | -34.0092 | -3.2531
1.7397 | -34.0094 | -3.2654
1.7458 | -34.0095 | -3.2777
1.7519 | -34.0096 | -3.2899
1.758 | -34.0097 | -3.3017
1.7641 | -34.0099 | -3.3133
1.7702 | -34.0099 | -3.3262
1.7763 | -34.0099 | -3.3376
1.7824 | -34.01 | -3.35
1.7885 | -34.01 | -3.3623
1.7947 | -34.0102 | -3.3742
1.8008 | -34.0102 | -3.3855
1.8069 | -34.0103 | -3.397
1.813 | -34.0104 | -3.4092
1.8191 | -34.0104 | -3.4218
1.8252 | -34.0106 | -3.4344
1.8313 | -34.0108 | -3.4465
1.8374 | -34.011 | -3.4583
1.8435 | -34.0111 | -3.4701
1.8496 | -34.0112 | -3.4829
1.8557 | -34.0114 | -3.4954
1.8618 | -34.0117 | -3.5075
1.8679 | -34.0118 | -3.5194
1.874 | -34.012 | -3.5313
1.8801 | -34.0122 | -3.5436
1.8862 | -34.0124 | -3.5555
1.8923 | -34.0126 | -3.5676
1.8984 | -34.0127 | -3.5795
1.9045 | -34.0129 | -3.591
1.9106 | -34.0129 | -3.6029
1.9167 | -34.013 | -3.6149
1.9228 | -34.0132 | -3.6267
1.9289 | -34.0134 | -3.6383
1.9351 | -34.0136 | -3.6489
1.9412 | -34.0138 | -3.6601
1.9473 | -34.0142 | -3.672
1.9534 | -34.0144 | -3.6832
1.9595 | -34.0146 | -3.6949
1.9656 | -34.015 | -3.707
1.9717 | -34.0152 | -3.7185
1.9778 | -34.0156 | -3.7297
1.9839 | -34.0159 | -3.741
1.99 | -34.0161 | -3.7519
1.9961 | -34.0162 | -3.7639
2.0022 | -34.0165 | -3.7751
2.0083 | -34.0167 | -3.7862
2.0144 | -34.0169 | -3.7974
2.0205 | -34.0173 | -3.809
2.0266 | -34.0177 | -3.8201
2.0327 | -34.018 | -3.8315
2.0388 | -34.0184 | -3.843
2.0449 | -34.0187 | -3.8542
2.051 | -34.019 | -3.8655
2.0571 | -34.0194 | -3.8766
2.0632 | -34.0196 | -3.8879
2.0693 | -34.0198 | -3.8999
2.0754 | -34.0202 | -3.9112
2.0816 | -34.0205 | -3.9232
2.0877 | -34.0208 | -3.9346
2.0938 | -34.0212 | -3.9465
2.0999 | -34.0214 | -3.9581
2.106 | -34.0217 | -3.9695
2.1121 | -34.0219 | -3.9813
2.1182 | -34.0222 | -3.992
2.1243 | -34.0225 | -4.0025
2.1304 | -34.0228 | -4.0136
2.1365 | -34.0231 | -4.0253
2.1426 | -34.0234 | -4.0373
2.1487 | -34.0238 | -4.0484
2.1548 | -34.0242 | -4.06
2.1609 | -34.0247 | -4.0713
2.167 | -34.0251 | -4.0825
2.1731 | -34.0255 | -4.0943
2.1792 | -34.0257 | -4.1054
2.1853 | -34.0262 | -4.1172
2.1914 | -34.0265 | -4.1285
2.1975 | -34.0269 | -4.14
2.2036 | -34.0272 | -4.1519
2.2097 | -34.0275 | -4.1633
2.2158 | -34.0279 | -4.1739
2.222 | -34.0282 | -4.1843
2.2281 | -34.0285 | -4.1956
2.2342 | -34.0288 | -4.2069
2.2403 | -34.0291 | -4.2175
2.2464 | -34.0294 | -4.2284
2.2525 | -34.0297 | -4.2397
2.2586 | -34.03 | -4.2502
2.2647 | -34.0304 | -4.2607
2.2708 | -34.0307 | -4.272
2.2769 | -34.0311 | -4.2826
2.283 | -34.0313 | -4.2938
2.2891 | -34.0315 | -4.3047
2.2952 | -34.0317 | -4.3154
2.3013 | -34.032 | -4.3263
2.3074 | -34.0322 | -4.3367
2.3135 | -34.0325 | -4.3482
2.3196 | -34.0326 | -4.3584
2.3257 | -34.0329 | -4.3691
2.3318 | -34.033 | -4.3802
2.3379 | -34.0332 | -4.391
2.344 | -34.0334 | -4.4024
2.3501 | -34.0336 | -4.4131
2.3562 | -34.0337 | -4.4236
2.3623 | -34.0339 | -4.434
2.3685 | -34.034 | -4.4447
2.3746 | -34.0341 | -4.4559
2.3807 | -34.0343 | -4.4665
2.3868 | -34.0345 | -4.4773
2.3929 | -34.0347 | -4.4881
2.399 | -34.0347 | -4.4985
2.4051 | -34.035 | -4.5091
2.4112 | -34.0352 | -4.5198
2.4173 | -34.0353 | -4.5305
2.4234 | -34.0354 | -4.5414
2.4295 | -34.0355 | -4.5523
2.4356 | -34.0357 | -4.5633
2.4417 | -34.0358 | -4.5744
2.4478 | -34.0359 | -4.5854
2.4539 | -34.036 | -4.5958
2.46 | -34.0361 | -4.6073
2.4661 | -34.0363 | -4.6184
2.4722 | -34.0364 | -4.6287
2.4783 | -34.0365 | -4.6392
2.4844 | -34.0365 | -4.6504
2.4905 | -34.0366 | -4.6613
2.4966 | -34.0368 | -4.6721
2.5027 | -34.0369 | -4.6826
2.5089 | -34.037 | -4.6936
2.515 | -34.037 | -4.7045
2.5211 | -34.0371 | -4.7153
2.5272 | -34.0371 | -4.7262
2.5333 | -34.0373 | -4.7372
2.5394 | -34.0374 | -4.748
2.5455 | -34.0375 | -4.7594
2.5516 | -34.0377 | -4.7698
2.5577 | -34.0378 | -4.7813
2.5638 | -34.038 | -4.793
2.5699 | -34.038 | -4.8045
2.576 | -34.0381 | -4.8161
2.5821 | -34.0382 | -4.8274
2.5882 | -34.0383 | -4.8385
2.5943 | -34.0384 | -4.8493
2.6004 | -34.0386 | -4.8604
2.6065 | -34.0387 | -4.8708
2.6126 | -34.0388 | -4.8812
2.6187 | -34.039 | -4.8917
2.6248 | -34.0391 | -4.9025
2.6309 | -34.0393 | -4.914
2.637 | -34.0396 | -4.9256
2.6431 | -34.0397 | -4.9361
2.6492 | -34.0399 | -4.9475
2.6554 | -34.0401 | -4.9583
2.6615 | -34.0402 | -4.9693
2.6676 | -34.0405 | -4.9807
2.6737 | -34.0408 | -4.9918
2.6798 | -34.0411 | -5.0029
2.6859 | -34.0413 | -5.0141
2.692 | -34.0415 | -5.0249
2.6981 | -34.0417 | -5.036
2.7042 | -34.0419 | -5.0471
2.7103 | -34.0421 | -5.0577
2.7164 | -34.0425 | -5.0686
2.7225 | -34.0428 | -5.0799
2.7286 | -34.043 | -5.0909
2.7347 | -34.0432 | -5.1014
2.7408 | -34.0436 | -5.1115
2.7469 | -34.0438 | -5.1226
2.753 | -34.0441 | -5.1332
2.7591 | -34.0445 | -5.1445
2.7652 | -34.0447 | -5.1559
2.7713 | -34.0449 | -5.1667
2.7774 | -34.0452 | -5.1775
2.7835 | -34.0454 | -5.1884
2.7896 | -34.0457 | -5.1989
2.7958 | -34.046 | -5.2096
2.8019 | -34.0463 | -5.2199
2.808 | -34.0466 | -5.2307
2.8141 | -34.047 | -5.2418
2.8202 | -34.0474 | -5.2526
2.8263 | -34.0477 | -5.2634
2.8324 | -34.0481 | -5.2748
2.8385 | -34.0485 | -5.2856
2.8446 | -34.0489 | -5.2963
2.8507 | -34.0493 | -5.3071
2.8568 | -34.0497 | -5.3175
2.8629 | -34.05 | -5.328
2.869 | -34.0504 | -5.339
2.8751 | -34.0508 | -5.3497
2.8812 | -34.0512 | -5.3609
2.8873 | -34.0516 | -5.3717
2.8934 | -34.0519 | -5.3826
2.8995 | -34.0523 | -5.3937
2.9056 | -34.0526 | -5.4048
2.9117 | -34.053 | -5.4157
2.9178 | -34.0534 | -5.4264
2.9239 | -34.0539 | -5.4379
2.93 | -34.0543 | -5.4492
2.9361 | -34.0546 | -5.46
2.9423 | -34.055 | -5.4706
2.9484 | -34.0555 | -5.4816
2.9545 | -34.0559 | -5.4926
2.9606 | -34.0564 | -5.5039
2.9667 | -34.0568 | -5.515
2.9728 | -34.0571 | -5.5262
2.9789 | -34.0576 | -5.5372
2.985 | -34.0581 | -5.5484
2.9911 | -34.0585 | -5.5597
2.9972 | -34.059 | -5.5707
3.0033 | -34.0594 | -5.5822
3.0094 | -34.0598 | -5.5932
3.0155 | -34.0603 | -5.6041
3.0216 | -34.0607 | -5.6153
3.0277 | -34.061 | -5.6269
3.0338 | -34.0614 | -5.6382
3.0399 | -34.0619 | -5.6495
3.046 | -34.0622 | -5.661
3.0521 | -34.0626 | -5.672
3.0582 | -34.0629 | -5.6839
3.0643 | -34.0633 | -5.6952
3.0704 | -34.0637 | -5.707
3.0765 | -34.064 | -5.718
3.0827 | -34.0643 | -5.7292
3.0888 | -34.0647 | -5.7404
3.0949 | -34.0651 | -5.7517
3.101 | -34.0655 | -5.7631
3.1071 | -34.0658 | -5.7748
3.1132 | -34.0662 | -5.7865
3.1193 | -34.0666 | -5.7979
3.1254 | -34.067 | -5.8092
3.1315 | -34.0674 | -5.821
3.1376 | -34.0677 | -5.8324
3.1437 | -34.0681 | -5.844
3.1498 | -34.0685 | -5.856
3.1559 | -34.0688 | -5.8674
3.162 | -34.0691 | -5.8794
3.1681 | -34.0694 | -5.8912
3.1742 | -34.0697 | -5.9027
3.1803 | -34.07 | -5.9143
3.1864 | -34.0704 | -5.9255
3.1925 | -34.0707 | -5.937
3.1986 | -34.071 | -5.9489
3.2047 | -34.0713 | -5.9608
3.2108 | -34.0716 | -5.9719
3.2169 | -34.0719 | -5.9835
3.223 | -34.0722 | -5.9951
3.2292 | -34.0725 | -6.0069
3.2353 | -34.0727 | -6.0177
3.2414 | -34.073 | -6.029
3.2475 | -34.0732 | -6.0404
3.2536 | -34.0735 | -6.0518
3.2597 | -34.0738 | -6.0626
3.2658 | -34.074 | -6.0736
3.2719 | -34.0742 | -6.0846
3.278 | -34.0745 | -6.0961
3.2841 | -34.0748 | -6.1071
3.2902 | -34.075 | -6.1185
3.2963 | -34.0753 | -6.1299
3.3024 | -34.0754 | -6.1409
3.3085 | -34.0758 | -6.1524
3.3146 | -34.0759 | -6.1636
3.3207 | -34.0762 | -6.175
3.3268 | -34.0764 | -6.1864
3.3329 | -34.0766 | -6.1971
3.339 | -34.0768 | -6.2078
3.3451 | -34.077 | -6.219
3.3512 | -34.0772 | -6.2297
3.3573 | -34.0774 | -6.2408
3.3634 | -34.0776 | -6.252
3.3696 | -34.0778 | -6.2633
3.3757 | -34.078 | -6.274
3.3818 | -34.0782 | -6.2848
3.3879 | -34.0784 | -6.2958
3.394 | -34.0786 | -6.3072
3.4001 | -34.0788 | -6.318
3.4062 | -34.0789 | -6.3288
3.4123 | -34.0791 | -6.3393
3.4184 | -34.0793 | -6.3506
3.4245 | -34.0795 | -6.3613
3.4306 | -34.0796 | -6.372
3.4367 | -34.0798 | -6.3825
3.4428 | -34.08 | -6.3928
3.4489 | -34.0802 | -6.4037
3.455 | -34.0803 | -6.4145
3.4611 | -34.0806 | -6.4257
3.4672 | -34.0807 | -6.4367
3.4733 | -34.0809 | -6.4477
3.4794 | -34.0811 | -6.4585
3.4855 | -34.0813 | -6.4688
3.4916 | -34.0815 | -6.4792
3.4977 | -34.0816 | -6.4904
3.5038 | -34.0818 | -6.5013
3.5099 | -34.0819 | -6.5115
3.5161 | -34.0821 | -6.5219
3.5222 | -34.0823 | -6.5322
3.5283 | -34.0824 | -6.543
3.5344 | -34.0825 | -6.5538
3.5405 | -34.0827 | -6.5643
3.5466 | -34.0829 | -6.5752
3.5527 | -34.083 | -6.586
3.5588 | -34.0831 | -6.5966
3.5649 | -34.0832 | -6.6072
3.571 | -34.0834 | -6.6178
3.5771 | -34.0835 | -6.6289
3.5832 | -34.0837 | -6.6387
3.5893 | -34.0839 | -6.6493
3.5954 | -34.0841 | -6.6601
3.6015 | -34.0843 | -6.6709
3.6076 | -34.0845 | -6.6814
3.6137 | -34.0847 | -6.6915
3.6198 | -34.085 | -6.7017
3.6259 | -34.0851 | -6.7126
3.632 | -34.0853 | -6.7232
3.6381 | -34.0855 | -6.7339
3.6442 | -34.0857 | -6.7445
3.6503 | -34.0859 | -6.7546
3.6565 | -34.0862 | -6.7654
3.6626 | -34.0863 | -6.7759
3.6687 | -34.0866 | -6.7865
3.6748 | -34.0868 | -6.7968
3.6809 | -34.0871 | -6.8075
3.687 | -34.0873 | -6.8184
3.6931 | -34.0876 | -6.8286
3.6992 | -34.0878 | -6.8393
3.7053 | -34.088 | -6.8498
3.7114 | -34.0881 | -6.8604
3.7175 | -34.0884 | -6.8711
3.7236 | -34.0887 | -6.881
3.7297 | -34.0889 | -6.8918
3.7358 | -34.0891 | -6.9027
3.7419 | -34.0894 | -6.9134
3.748 | -34.0896 | -6.9239
3.7541 | -34.0898 | -6.9344
3.7602 | -34.0901 | -6.945
3.7663 | -34.0903 | -6.9554
3.7724 | -34.0906 | -6.966
3.7785 | -34.0909 | -6.9764
3.7846 | -34.0911 | -6.9869
3.7907 | -34.0914 | -6.9978
3.7969 | -34.0917 | -7.0085
3.803 | -34.092 | -7.0186
3.8091 | -34.0922 | -7.0292
3.8152 | -34.0925 | -7.0397
3.8213 | -34.0928 | -7.0506
3.8274 | -34.0931 | -7.0608
3.8335 | -34.0934 | -7.0716
3.8396 | -34.0937 | -7.0822
3.8457 | -34.094 | -7.093
3.8518 | -34.0944 | -7.1034
3.8579 | -34.0948 | -7.1139
3.864 | -34.095 | -7.1247
3.8701 | -34.0954 | -7.1356
3.8762 | -34.0957 | -7.146
3.8823 | -34.0959 | -7.1568
3.8884 | -34.0963 | -7.1674
3.8945 | -34.0967 | -7.1783
3.9006 | -34.097 | -7.1892
3.9067 | -34.0973 | -7.2003
3.9128 | -34.0976 | -7.211
3.9189 | -34.0979 | -7.2215
3.925 | -34.0981 | -7.2323
3.9311 | -34.0984 | -7.2425
3.9372 | -34.0988 | -7.2533
3.9434 | -34.0991 | -7.264
3.9495 | -34.0995 | -7.2747
3.9556 | -34.0998 | -7.2852
3.9617 | -34.1001 | -7.2954
3.9678 | -34.1005 | -7.3059
3.9739 | -34.1008 | -7.3165
3.98 | -34.1011 | -7.3272
3.9861 | -34.1014 | -7.3379
3.9922 | -34.1016 | -7.3485
3.9983 | -34.1019 | -7.3587
4.0044 | -34.1022 | -7.3689
4.0105 | -34.1026 | -7.3796
4.0166 | -34.1029 | -7.3903
4.0227 | -34.1032 | -7.4009
4.0288 | -34.1035 | -7.4119
4.0349 | -34.1038 | -7.4223
4.041 | -34.1041 | -7.433
4.0471 | -34.1043 | -7.4437
4.0532 | -34.1046 | -7.4539
4.0593 | -34.105 | -7.4647
4.0654 | -34.1053 | -7.4751
4.0715 | -34.1056 | -7.4863
4.0776 | -34.1059 | -7.4968
4.0838 | -34.1063 | -7.5079
4.0899 | -34.1066 | -7.5182
4.096 | -34.1068 | -7.5289
4.1021 | -34.1071 | -7.5392
4.1082 | -34.1074 | -7.5496
4.1143 | -34.1077 | -7.5598
4.1204 | -34.108 | -7.5702
4.1265 | -34.1083 | -7.5805
4.1326 | -34.1086 | -7.5908
4.1387 | -34.109 | -7.6012
4.1448 | -34.1092 | -7.6119
4.1509 | -34.1095 | -7.6225
4.157 | -34.1097 | -7.6335
4.1631 | -34.11 | -7.6442
4.1692 | -34.1103 | -7.6546
4.1753 | -34.1107 | -7.6652
4.1814 | -34.1109 | -7.6754
4.1875 | -34.1112 | -7.6862
4.1936 | -34.1115 | -7.6968
4.1997 | -34.1117 | -7.707
4.2058 | -34.112 | -7.7178
4.2119 | -34.1123 | -7.7281
4.218 | -34.1125 | -7.7384
4.2241 | -34.1128 | -7.7487
4.2303 | -34.1131 | -7.759
4.2364 | -34.1133 | -7.7694
4.2425 | -34.1135 | -7.78
4.2486 | -34.1138 | -7.7906
4.2547 | -34.114 | -7.8009
4.2608 | -34.1142 | -7.8115
4.2669 | -34.1144 | -7.8217
4.273 | -34.1147 | -7.8322
4.2791 | -34.115 | -7.8426
4.2852 | -34.1152 | -7.8534
4.2913 | -34.1154 | -7.8636
4.2974 | -34.1157 | -7.8743
4.3035 | -34.1159 | -7.8851
4.3096 | -34.1161 | -7.8957
4.3157 | -34.1164 | -7.906
4.3218 | -34.1166 | -7.9163
4.3279 | -34.1169 | -7.9268
4.334 | -34.1171 | -7.9371
4.3401 | -34.1173 | -7.9472
4.3462 | -34.1176 | -7.9578
4.3523 | -34.1179 | -7.9678
4.3584 | -34.1181 | -7.9778
4.3645 | -34.1183 | -7.9881
4.3707 | -34.1185 | -7.9982
4.3768 | -34.1188 | -8.0086
4.3829 | -34.119 | -8.0193
4.389 | -34.1192 | -8.0296
4.3951 | -34.1195 | -8.0398
4.4012 | -34.1197 | -8.0504
4.4073 | -34.1199 | -8.0606
4.4134 | -34.1201 | -8.0709
4.4195 | -34.1204 | -8.0812
4.4256 | -34.1207 | -8.0913
4.4317 | -34.1209 | -8.1016
4.4378 | -34.1211 | -8.1116
4.4439 | -34.1213 | -8.1216
4.45 | -34.1215 | -8.1319
4.4561 | -34.1217 | -8.1422
4.4622 | -34.1219 | -8.1525
4.4683 | -34.1221 | -8.1628
4.4744 | -34.1223 | -8.1734
4.4805 | -34.1226 | -8.1833
4.4866 | -34.1228 | -8.1935
4.4927 | -34.1231 | -8.204
4.4988 | -34.1233 | -8.2143
4.5049 | -34.1235 | -8.2242
4.511 | -34.1238 | -8.2342
4.5172 | -34.124 | -8.2445
4.5233 | -34.1242 | -8.2551
4.5294 | -34.1245 | -8.2652
4.5355 | -34.1247 | -8.2755
4.5416 | -34.1249 | -8.2857
4.5477 | -34.1252 | -8.2959
4.5538 | -34.1254 | -8.3058
4.5599 | -34.1257 | -8.316
4.566 | -34.1259 | -8.3263
4.5721 | -34.1261 | -8.3365
4.5782 | -34.1264 | -8.347
4.5843 | -34.1266 | -8.3573
4.5904 | -34.1269 | -8.3677
4.5965 | -34.1271 | -8.3778
4.6026 | -34.1274 | -8.3884
4.6087 | -34.1276 | -8.3986
4.6148 | -34.1278 | -8.4092
4.6209 | -34.1281 | -8.4194
4.627 | -34.1283 | -8.4298
4.6331 | -34.1285 | -8.4399
4.6392 | -34.1287 | -8.4503
4.6453 | -34.129 | -8.4605
4.6514 | -34.1293 | -8.4709
4.6576 | -34.1295 | -8.4815
4.6637 | -34.1297 | -8.4914
4.6698 | -34.1299 | -8.5015
4.6759 | -34.1301 | -8.5122
4.682 | -34.1303 | -8.5224
4.6881 | -34.1305 | -8.5327
4.6942 | -34.1308 | -8.5429
4.7003 | -34.131 | -8.5532
4.7064 | -34.1312 | -8.5633
4.7125 | -34.1315 | -8.5739
4.7186 | -34.1317 | -8.584
4.7247 | -34.132 | -8.5944
4.7308 | -34.1322 | -8.6046
4.7369 | -34.1325 | -8.615
4.743 | -34.1327 | -8.6257
4.7491 | -34.1329 | -8.6363
4.7552 | -34.1332 | -8.6467
4.7613 | -34.1334 | -8.6571
4.7674 | -34.1336 | -8.6673
4.7735 | -34.1338 | -8.678
4.7796 | -34.1341 | -8.6886
4.7857 | -34.1344 | -8.6991
4.7918 | -34.1346 | -8.7092
4.7979 | -34.1349 | -8.7194
4.8041 | -34.1351 | -8.7298
4.8102 | -34.1354 | -8.7398
4.8163 | -34.1356 | -8.7501
4.8224 | -34.1359 | -8.7604
4.8285 | -34.1361 | -8.7705
4.8346 | -34.1364 | -8.7806
4.8407 | -34.1366 | -8.7911
4.8468 | -34.1369 | -8.8017
4.8529 | -34.1372 | -8.8123
4.859 | -34.1374 | -8.8229
4.8651 | -34.1377 | -8.8331
4.8712 | -34.1379 | -8.8433
4.8773 | -34.1381 | -8.8537
4.8834 | -34.1384 | -8.8641
4.8895 | -34.1386 | -8.8742
4.8956 | -34.1388 | -8.8845
4.9017 | -34.139 | -8.8947
4.9078 | -34.1393 | -8.9054
4.9139 | -34.1395 | -8.9154
4.92 | -34.1399 | -8.9254
4.9261 | -34.1401 | -8.9354
4.9322 | -34.1404 | -8.9457
4.9383 | -34.1407 | -8.956
4.9445 | -34.1409 | -8.9664
4.9506 | -34.1412 | -8.9766
4.9567 | -34.1414 | -8.9867
4.9628 | -34.1417 | -8.997
4.9689 | -34.1419 | -9.0077
4.975 | -34.1422 | -9.0183
4.9811 | -34.1425 | -9.0283
4.9872 | -34.1428 | -9.0385
4.9933 | -34.1431 | -9.0491
4.9994 | -34.1434 | -9.0598
5.0055 | -34.1437 | -9.0702
5.0116 | -34.144 | -9.0809
5.0177 | -34.1443 | -9.0915
5.0238 | -34.1446 | -9.1022
5.0299 | -34.1449 | -9.1128
5.036 | -34.1452 | -9.1231
5.0421 | -34.1455 | -9.1334
5.0482 | -34.1458 | -9.1438
5.0543 | -34.1461 | -9.1539
5.0604 | -34.1464 | -9.1642
5.0665 | -34.1467 | -9.1743
5.0726 | -34.147 | -9.1849
5.0787 | -34.1473 | -9.1955
5.0848 | -34.1476 | -9.206
5.091 | -34.1479 | -9.2161
5.0971 | -34.1482 | -9.2261
5.1032 | -34.1485 | -9.2369
5.1093 | -34.1488 | -9.2472
5.1154 | -34.149 | -9.2573
5.1215 | -34.1493 | -9.2675
5.1276 | -34.1496 | -9.2777
5.1337 | -34.15 | -9.2878
5.1398 | -34.1502 | -9.2979
5.1459 | -34.1506 | -9.3085
5.152 | -34.1509 | -9.319
5.1581 | -34.1511 | -9.3295
5.1642 | -34.1515 | -9.3399
5.1703 | -34.1518 | -9.3501
5.1764 | -34.1521 | -9.3603
5.1825 | -34.1524 | -9.3704
5.1886 | -34.1527 | -9.3804
5.1947 | -34.153 | -9.3908
5.2008 | -34.1534 | -9.4012
5.2069 | -34.1536 | -9.4117
5.213 | -34.154 | -9.4218
5.2191 | -34.1543 | -9.4322
5.2252 | -34.1546 | -9.4426
5.2314 | -34.1549 | -9.4531
5.2375 | -34.1552 | -9.4631
5.2436 | -34.1555 | -9.4735
5.2497 | -34.1558 | -9.4835
5.2558 | -34.1561 | -9.4942
5.2619 | -34.1565 | -9.5048
5.268 | -34.1568 | -9.5153
5.2741 | -34.1572 | -9.5256
5.2802 | -34.1575 | -9.5363
5.2863 | -34.1579 | -9.5466
5.2924 | -34.1582 | -9.5571
5.2985 | -34.1585 | -9.5675
5.3046 | -34.1588 | -9.5779
5.3107 | -34.1592 | -9.5884
5.3168 | -34.1595 | -9.5988
5.3229 | -34.1598 | -9.6092
5.329 | -34.1601 | -9.6196
5.3351 | -34.1605 | -9.6296
5.3412 | -34.1608 | -9.6399
5.3473 | -34.1611 | -9.6503
5.3534 | -34.1615 | -9.6608
5.3595 | -34.1618 | -9.6711
5.3656 | -34.1621 | -9.6814
5.3717 | -34.1624 | -9.6921
5.3779 | -34.1628 | -9.7028
5.384 | -34.1631 | -9.7131
5.3901 | -34.1635 | -9.7237
5.3962 | -34.1638 | -9.7339
5.4023 | -34.1642 | -9.7445
5.4084 | -34.1645 | -9.7547
5.4145 | -34.1649 | -9.7648
5.4206 | -34.1652 | -9.7753
5.4267 | -34.1655 | -9.7859
5.4328 | -34.1659 | -9.7963
5.4389 | -34.1662 | -9.8064
5.445 | -34.1665 | -9.8171
5.4511 | -34.1669 | -9.8273
5.4572 | -34.1672 | -9.8378
5.4633 | -34.1675 | -9.8481
5.4694 | -34.1678 | -9.8585
5.4755 | -34.1682 | -9.8686
5.4816 | -34.1685 | -9.8791
5.4877 | -34.1689 | -9.8893
5.4938 | -34.1692 | -9.8995
5.4999 | -34.1696 | -9.9097
5.506 | -34.1699 | -9.92
5.5121 | -34.1702 | -9.9303
5.5183 | -34.1706 | -9.9404
5.5244 | -34.1709 | -9.951
5.5305 | -34.1712 | -9.9612
5.5366 | -34.1716 | -9.9713
5.5427 | -34.1719 | -9.9818
5.5488 | -34.1722 | -9.9922
5.5549 | -34.1725 | -10.0026
5.561 | -34.1729 | -10.0131
5.5671 | -34.1732 | -10.0233
5.5732 | -34.1735 | -10.0339
5.5793 | -34.1739 | -10.0445
5.5854 | -34.1743 | -10.0548
5.5915 | -34.1746 | -10.0655
5.5976 | -34.1749 | -10.0759
5.6037 | -34.1752 | -10.0863
5.6098 | -34.1755 | -10.0966
5.6159 | -34.1759 | -10.1067
5.622 | -34.1762 | -10.1172
5.6281 | -34.1766 | -10.1276
5.6342 | -34.1769 | -10.1383
5.6403 | -34.1772 | -10.1487
5.6464 | -34.1776 | -10.1592
5.6525 | -34.1779 | -10.1697
5.6586 | -34.1782 | -10.1797
5.6648 | -34.1786 | -10.1901
5.6709 | -34.179 | -10.2007
5.677 | -34.1793 | -10.211
5.6831 | -34.1796 | -10.2215
5.6892 | -34.1799 | -10.232
5.6953 | -34.1803 | -10.2424
5.7014 | -34.1806 | -10.2528
5.7075 | -34.1809 | -10.2632
5.7136 | -34.1812 | -10.2735
5.7197 | -34.1815 | -10.2839
5.7258 | -34.1818 | -10.294
5.7319 | -34.1822 | -10.3044
5.738 | -34.1825 | -10.3148
5.7441 | -34.1829 | -10.3251
5.7502 | -34.1832 | -10.3352
5.7563 | -34.1835 | -10.3456
5.7624 | -34.1839 | -10.3561
5.7685 | -34.1842 | -10.3666
5.7746 | -34.1845 | -10.3772
5.7807 | -34.1848 | -10.3876
5.7868 | -34.1851 | -10.3982
5.7929 | -34.1854 | -10.4084
5.799 | -34.1858 | -10.4186
5.8052 | -34.1861 | -10.4288
5.8113 | -34.1864 | -10.4391
5.8174 | -34.1867 | -10.4493
5.8235 | -34.187 | -10.4597
5.8296 | -34.1873 | -10.4702
5.8357 | -34.1876 | -10.4804
5.8418 | -34.188 | -10.4909
5.8479 | -34.1883 | -10.5013
5.854 | -34.1886 | -10.5118
5.8601 | -34.1889 | -10.5221
5.8662 | -34.1892 | -10.5326
5.8723 | -34.1895 | -10.5428
5.8784 | -34.1898 | -10.553
5.8845 | -34.1902 | -10.5633
5.8906 | -34.1905 | -10.5737
5.8967 | -34.1908 | -10.5839
5.9028 | -34.1911 | -10.594
5.9089 | -34.1915 | -10.6043
5.915 | -34.1918 | -10.6144
5.9211 | -34.1921 | -10.6246
5.9272 | -34.1925 | -10.6347
5.9333 | -34.1928 | -10.6449
5.9394 | -34.1931 | -10.6552
5.9455 | -34.1934 | -10.6653
5.9517 | -34.1938 | -10.6755
5.9578 | -34.1942 | -10.6859
5.9639 | -34.1945 | -10.6959
5.97 | -34.1949 | -10.7063
5.9761 | -34.1952 | -10.7165
5.9822 | -34.1955 | -10.7269
5.9883 | -34.1958 | -10.7372
5.9944 | -34.1962 | -10.7476
6.0005 | -34.1965 | -10.7581
6.0066 | -34.1968 | -10.7682
6.0127 | -34.1972 | -10.7782
6.0188 | -34.1975 | -10.7887
6.0249 | -34.1978 | -10.7991
6.031 | -34.1981 | -10.8092
6.0371 | -34.1984 | -10.8194
6.0432 | -34.1988 | -10.8296
6.0493 | -34.1991 | -10.8398
6.0554 | -34.1994 | -10.8499
6.0615 | -34.1998 | -10.8602
6.0676 | -34.2001 | -10.8706
6.0737 | -34.2005 | -10.881
6.0798 | -34.2007 | -10.8913
6.0859 | -34.2011 | -10.9015
6.0921 | -34.2014 | -10.9115
6.0982 | -34.2017 | -10.9217
6.1043 | -34.2021 | -10.9319
6.1104 | -34.2024 | -10.9423
6.1165 | -34.2028 | -10.9522
6.1226 | -34.2031 | -10.9623
6.1287 | -34.2034 | -10.9725
6.1348 | -34.2038 | -10.9827
6.1409 | -34.2041 | -10.9929
6.147 | -34.2044 | -11.0031
6.1531 | -34.2047 | -11.0133
6.1592 | -34.2051 | -11.0234
6.1653 | -34.2054 | -11.0335
6.1714 | -34.2057 | -11.0438
6.1775 | -34.206 | -11.0541
6.1836 | -34.2063 | -11.0643
6.1897 | -34.2067 | -11.0746
6.1958 | -34.207 | -11.0851
6.2019 | -34.2074 | -11.0954
6.208 | -34.2077 | -11.1056
6.2141 | -34.208 | -11.1158
6.2202 | -34.2084 | -11.1261
6.2263 | -34.2087 | -11.1363
6.2325 | -34.209 | -11.1462
6.2386 | -34.2094 | -11.1561
6.2447 | -34.2097 | -11.1666
6.2508 | -34.2101 | -11.1768
6.2569 | -34.2104 | -11.1869
6.263 | -34.2107 | -11.1971
6.2691 | -34.211 | -11.2072
6.2752 | -34.2113 | -11.2175
6.2813 | -34.2117 | -11.2277
6.2874 | -34.212 | -11.2378
6.2935 | -34.2123 | -11.2479
6.2996 | -34.2126 | -11.258
6.3057 | -34.213 | -11.2683
6.3118 | -34.2132 | -11.2786
6.3179 | -34.2136 | -11.2888
6.324 | -34.2139 | -11.299
6.3301 | -34.2142 | -11.3092
6.3362 | -34.2146 | -11.3193
6.3423 | -34.2149 | -11.3296
6.3484 | -34.2152 | -11.3399
6.3545 | -34.2156 | -11.3499
6.3606 | -34.2159 | -11.3598
6.3667 | -34.2162 | -11.3698
6.3728 | -34.2165 | -11.38
6.379 | -34.2168 | -11.3902
6.3851 | -34.2172 | -11.4001
6.3912 | -34.2175 | -11.4103
6.3973 | -34.2178 | -11.4205
6.4034 | -34.2181 | -11.4305
6.4095 | -34.2184 | -11.4407
6.4156 | -34.2187 | -11.4509
6.4217 | -34.219 | -11.461
6.4278 | -34.2194 | -11.4712
6.4339 | -34.2197 | -11.4812
6.44 | -34.2201 | -11.4913
6.4461 | -34.2204 | -11.5016
6.4522 | -34.2207 | -11.5115
6.4583 | -34.221 | -11.5218
6.4644 | -34.2213 | -11.532
6.4705 | -34.2216 | -11.5419
6.4766 | -34.222 | -11.5521
6.4827 | -34.2223 | -11.5623
6.4888 | -34.2226 | -11.5722
6.4949 | -34.2229 | -11.5821
6.501 | -34.2232 | -11.5923
6.5071 | -34.2236 | -11.6023
6.5132 | -34.2239 | -11.6125
6.5194 | -34.2242 | -11.6227
6.5255 | -34.2245 | -11.6326
6.5316 | -34.2248 | -11.6428
6.5377 | -34.2252 | -11.6529
6.5438 | -34.2255 | -11.6629
6.5499 | -34.2258 | -11.6729
6.556 | -34.2261 | -11.683
6.5621 | -34.2264 | -11.6933
6.5682 | -34.2268 | -11.7035
6.5743 | -34.2271 | -11.7136
6.5804 | -34.2274 | -11.7238
6.5865 | -34.2278 | -11.7339
6.5926 | -34.2281 | -11.7439
6.5987 | -34.2284 | -11.7538
6.6048 | -34.2287 | -11.764
6.6109 | -34.2291 | -11.7742
6.617 | -34.2294 | -11.7844
6.6231 | -34.2297 | -11.7945
6.6292 | -34.23 | -11.8046
6.6353 | -34.2303 | -11.815
6.6414 | -34.2307 | -11.8251
6.6475 | -34.231 | -11.8352
6.6536 | -34.2313 | -11.8453
6.6597 | -34.2316 | -11.8551
6.6659 | -34.232 | -11.8652
6.672 | -34.2323 | -11.8753
6.6781 | -34.2326 | -11.8854
6.6842 | -34.233 | -11.8953
6.6903 | -34.2333 | -11.9053
6.6964 | -34.2337 | -11.9154
6.7025 | -34.234 | -11.9253
6.7086 | -34.2343 | -11.9354
6.7147 | -34.2346 | -11.9456
6.7208 | -34.2349 | -11.9558
6.7269 | -34.2352 | -11.9658
6.733 | -34.2355 | -11.9757
6.7391 | -34.2359 | -11.9858
6.7452 | -34.2362 | -11.9957
6.7513 | -34.2365 | -12.0055
6.7574 | -34.2369 | -12.0156
6.7635 | -34.2372 | -12.0255
6.7696 | -34.2375 | -12.0355
6.7757 | -34.2379 | -12.0455
6.7818 | -34.2382 | -12.0556
6.7879 | -34.2386 | -12.0656
6.794 | -34.2389 | -12.0755
6.8001 | -34.2392 | -12.0855
6.8063 | -34.2395 | -12.0955
6.8124 | -34.2398 | -12.1058
6.8185 | -34.2402 | -12.116
6.8246 | -34.2405 | -12.126
6.8307 | -34.2408 | -12.1359
6.8368 | -34.2411 | -12.1462
6.8429 | -34.2414 | -12.1563
6.849 | -34.2418 | -12.1664
6.8551 | -34.2421 | -12.1763
6.8612 | -34.2425 | -12.1863
6.8673 | -34.2428 | -12.1963
6.8734 | -34.2431 | -12.2065
6.8795 | -34.2434 | -12.2166
6.8856 | -34.2438 | -12.2266
6.8917 | -34.2441 | -12.2366
6.8978 | -34.2444 | -12.2465
6.9039 | -34.2448 | -12.2562
6.91 | -34.2451 | -12.2663
6.9161 | -34.2455 | -12.2762
6.9222 | -34.2458 | -12.2861
6.9283 | -34.2462 | -12.2959
6.9344 | -34.2465 | -12.3059
6.9405 | -34.2469 | -12.3158
6.9466 | -34.2472 | -12.326
6.9528 | -34.2476 | -12.3362
6.9589 | -34.2479 | -12.3464
6.965 | -34.2483 | -12.3565
6.9711 | -34.2486 | -12.3666
6.9772 | -34.249 | -12.3766
6.9833 | -34.2493 | -12.3865
6.9894 | -34.2497 | -12.3968
6.9955 | -34.25 | -12.4066
7.0016 | -34.2504 | -12.4166
7.0077 | -34.2507 | -12.4266
7.0138 | -34.2511 | -12.4365
7.0199 | -34.2514 | -12.4465
7.026 | -34.2517 | -12.4564
7.0321 | -34.2521 | -12.4662
7.0382 | -34.2524 | -12.4759
7.0443 | -34.2528 | -12.486
7.0504 | -34.2531 | -12.4961
7.0565 | -34.2535 | -12.5061
7.0626 | -34.2538 | -12.5158
7.0687 | -34.2542 | -12.5254
7.0748 | -34.2545 | -12.5351
7.0809 | -34.2549 | -12.545
7.087 | -34.2552 | -12.5552
7.0932 | -34.2555 | -12.5652
7.0993 | -34.2559 | -12.5751
7.1054 | -34.2562 | -12.5849
7.1115 | -34.2566 | -12.5947
7.1176 | -34.2569 | -12.6045
7.1237 | -34.2572 | -12.6145
7.1298 | -34.2576 | -12.6243
7.1359 | -34.2579 | -12.6343
7.142 | -34.2583 | -12.6439
7.1481 | -34.2586 | -12.6537
7.1542 | -34.2589 | -12.6637
7.1603 | -34.2593 | -12.6736
7.1664 | -34.2597 | -12.6835
7.1725 | -34.26 | -12.6934
7.1786 | -34.2604 | -12.7033
7.1847 | -34.2607 | -12.7132
7.1908 | -34.2611 | -12.7233
7.1969 | -34.2614 | -12.7333
7.203 | -34.2618 | -12.7431
7.2091 | -34.2621 | -12.7528
7.2152 | -34.2625 | -12.7627
7.2213 | -34.2628 | -12.7724
7.2274 | -34.2632 | -12.7823
7.2335 | -34.2635 | -12.7921
7.2397 | -34.2639 | -12.8019
7.2458 | -34.2643 | -12.8115
7.2519 | -34.2647 | -12.8214
7.258 | -34.265 | -12.8315
7.2641 | -34.2654 | -12.8414
7.2702 | -34.2657 | -12.8514
7.2763 | -34.2661 | -12.8613
7.2824 | -34.2664 | -12.8713
7.2885 | -34.2668 | -12.8812
7.2946 | -34.2671 | -12.891
7.3007 | -34.2675 | -12.901
7.3068 | -34.2678 | -12.9111
7.3129 | -34.2682 | -12.921
7.319 | -34.2685 | -12.9308
7.3251 | -34.2689 | -12.9405
7.3312 | -34.2692 | -12.9504
7.3373 | -34.2696 | -12.9601
7.3434 | -34.27 | -12.9698
7.3495 | -34.2704 | -12.9795
7.3556 | -34.2707 | -12.9895
7.3617 | -34.2711 | -12.9995
7.3678 | -34.2714 | -13.0093
7.3739 | -34.2718 | -13.0193
7.3801 | -34.2721 | -13.0294
7.3862 | -34.2725 | -13.0391
7.3923 | -34.2729 | -13.0492
7.3984 | -34.2732 | -13.0588
7.4045 | -34.2736 | -13.0686
7.4106 | -34.274 | -13.0784
7.4167 | -34.2743 | -13.0884
7.4228 | -34.2747 | -13.0983
7.4289 | -34.2751 | -13.1081
7.435 | -34.2754 | -13.1177
7.4411 | -34.2758 | -13.1276
7.4472 | -34.2761 | -13.1375
7.4533 | -34.2765 | -13.1471
7.4594 | -34.2769 | -13.1571
7.4655 | -34.2772 | -13.1668
7.4716 | -34.2776 | -13.1769
7.4777 | -34.278 | -13.1867
7.4838 | -34.2783 | -13.1963
7.4899 | -34.2787 | -13.206
7.496 | -34.279 | -13.2156
7.5021 | -34.2794 | -13.2254
7.5082 | -34.2798 | -13.2353
7.5143 | -34.2801 | -13.2451
7.5204 | -34.2805 | -13.2551
7.5266 | -34.2808 | -13.2647
7.5327 | -34.2812 | -13.2744
7.5388 | -34.2815 | -13.2845
7.5449 | -34.2819 | -13.2945
7.551 | -34.2823 | -13.3042
7.5571 | -34.2827 | -13.3141
7.5632 | -34.283 | -13.3239
7.5693 | -34.2834 | -13.3336
7.5754 | -34.2838 | -13.3433
7.5815 | -34.2841 | -13.3532
7.5876 | -34.2845 | -13.3632
7.5937 | -34.2849 | -13.3731
7.5998 | -34.2853 | -13.3829
7.6059 | -34.2857 | -13.3929
7.612 | -34.286 | -13.4028
7.6181 | -34.2864 | -13.4125
7.6242 | -34.2867 | -13.4226
7.6303 | -34.2871 | -13.4325
7.6364 | -34.2875 | -13.4424
7.6425 | -34.2878 | -13.4523
7.6486 | -34.2882 | -13.4622
7.6547 | -34.2886 | -13.4719
7.6608 | -34.2889 | -13.4817
7.667 | -34.2893 | -13.4916
7.6731 | -34.2896 | -13.5015
7.6792 | -34.29 | -13.5115
7.6853 | -34.2904 | -13.5215
7.6914 | -34.2907 | -13.5314
7.6975 | -34.2911 | -13.5413
7.7036 | -34.2915 | -13.5512
7.7097 | -34.2919 | -13.5611
7.7158 | -34.2923 | -13.571
7.7219 | -34.2926 | -13.5807
7.728 | -34.293 | -13.5906
7.7341 | -34.2934 | -13.6004
7.7402 | -34.2937 | -13.6103
7.7463 | -34.2941 | -13.6202
7.7524 | -34.2945 | -13.6303
7.7585 | -34.2949 | -13.6402
7.7646 | -34.2952 | -13.65
7.7707 | -34.2956 | -13.66
7.7768 | -34.2959 | -13.67
7.7829 | -34.2963 | -13.6801
7.789 | -34.2967 | -13.6902
7.7951 | -34.297 | -13.7
7.8012 | -34.2974 | -13.7099
7.8073 | -34.2978 | -13.7199
7.8135 | -34.2982 | -13.7298
7.8196 | -34.2986 | -13.7398
7.8257 | -34.2989 | -13.7499
7.8318 | -34.2993 | -13.7598
7.8379 | -34.2997 | -13.7696
7.844 | -34.3 | -13.7794
7.8501 | -34.3004 | -13.7895
7.8562 | -34.3007 | -13.7996
7.8623 | -34.3012 | -13.8094
7.8684 | -34.3015 | -13.8193
7.8745 | -34.3018 | -13.8293
7.8806 | -34.3022 | -13.8393
7.8867 | -34.3026 | -13.849
7.8928 | -34.303 | -13.8587
7.8989 | -34.3034 | -13.8686
7.905 | -34.3037 | -13.8786
7.9111 | -34.3041 | -13.8885
7.9172 | -34.3044 | -13.8984
7.9233 | -34.3048 | -13.9084
7.9294 | -34.3052 | -13.9184
7.9355 | -34.3055 | -13.9281
7.9416 | -34.3059 | -13.9381
7.9477 | -34.3062 | -13.9478
7.9539 | -34.3066 | -13.9576
7.96 | -34.307 | -13.9676
7.9661 | -34.3073 | -13.9775
7.9722 | -34.3077 | -13.9872
7.9783 | -34.3081 | -13.9974
7.9844 | -34.3085 | -14.0072
7.9905 | -34.3088 | -14.0173
7.9966 | -34.3092 | -14.0273
8.0027 | -34.3096 | -14.0372
8.0088 | -34.3099 | -14.047
8.0149 | -34.3103 | -14.057
8.021 | -34.3106 | -14.0669
8.0271 | -34.311 | -14.077
8.0332 | -34.3114 | -14.0868
8.0393 | -34.3118 | -14.0968
8.0454 | -34.3121 | -14.107
8.0515 | -34.3126 | -14.1169
8.0576 | -34.3129 | -14.127
8.0637 | -34.3133 | -14.1368
8.0698 | -34.3136 | -14.1466
8.0759 | -34.314 | -14.1565
8.082 | -34.3143 | -14.1663
8.0881 | -34.3147 | -14.1762
8.0942 | -34.3151 | -14.1862
8.1004 | -34.3155 | -14.1962
8.1065 | -34.3159 | -14.206
8.1126 | -34.3162 | -14.216
8.1187 | -34.3166 | -14.226
8.1248 | -34.3169 | -14.2357
8.1309 | -34.3173 | -14.2454
8.137 | -34.3177 | -14.2551
8.1431 | -34.3181 | -14.2653
8.1492 | -34.3184 | -14.2752
8.1553 | -34.3188 | -14.2851
8.1614 | -34.3192 | -14.2952
8.1675 | -34.3195 | -14.3051
8.1736 | -34.3199 | -14.3152
8.1797 | -34.3202 | -14.3251
8.1858 | -34.3206 | -14.3351
8.1919 | -34.321 | -14.345
8.198 | -34.3213 | -14.3548
8.2041 | -34.3217 | -14.3647
8.2102 | -34.3221 | -14.3747
8.2163 | -34.3224 | -14.3844
8.2224 | -34.3228 | -14.3943
8.2285 | -34.3232 | -14.404
8.2346 | -34.3235 | -14.4139
8.2408 | -34.3239 | -14.4236
8.2469 | -34.3243 | -14.4336
8.253 | -34.3247 | -14.4434
8.2591 | -34.3251 | -14.4534
8.2652 | -34.3255 | -14.4633
8.2713 | -34.3258 | -14.4733
8.2774 | -34.3262 | -14.4833
8.2835 | -34.3266 | -14.4934
8.2896 | -34.3269 | -14.5031
8.2957 | -34.3273 | -14.5131
8.3018 | -34.3277 | -14.523
8.3079 | -34.3281 | -14.5328
8.314 | -34.3284 | -14.5427
8.3201 | -34.3288 | -14.5526
8.3262 | -34.3292 | -14.5625
8.3323 | -34.3296 | -14.5724
8.3384 | -34.3299 | -14.5824
8.3445 | -34.3303 | -14.5921
8.3506 | -34.3307 | -14.6021
8.3567 | -34.3311 | -14.6121
8.3628 | -34.3314 | -14.622
8.3689 | -34.3318 | -14.6318
8.375 | -34.3322 | -14.6417
8.3812 | -34.3325 | -14.6514
8.3873 | -34.3329 | -14.6613
8.3934 | -34.3333 | -14.6713
8.3995 | -34.3336 | -14.681
8.4056 | -34.334 | -14.6909
8.4117 | -34.3344 | -14.7008
8.4178 | -34.3347 | -14.7107
8.4239 | -34.3351 | -14.7205
8.43 | -34.3355 | -14.7304
8.4361 | -34.3358 | -14.7404
8.4422 | -34.3362 | -14.7502
8.4483 | -34.3366 | -14.7599
8.4544 | -34.337 | -14.7699
8.4605 | -34.3373 | -14.7796
8.4666 | -34.3377 | -14.7896
8.4727 | -34.338 | -14.7995
8.4788 | -34.3384 | -14.8093
8.4849 | -34.3388 | -14.8192
8.491 | -34.3391 | -14.8291
8.4971 | -34.3395 | -14.8388
8.5032 | -34.3399 | -14.8485
8.5093 | -34.3403 | -14.8582
8.5154 | -34.3407 | -14.8682
8.5215 | -34.341 | -14.8779
8.5277 | -34.3414 | -14.8878
8.5338 | -34.3418 | -14.8976
8.5399 | -34.3421 | -14.9074
8.546 | -34.3425 | -14.9171
8.5521 | -34.3429 | -14.927
8.5582 | -34.3433 | -14.9368
8.5643 | -34.3436 | -14.9464
8.5704 | -34.344 | -14.9561
8.5765 | -34.3444 | -14.966
8.5826 | -34.3448 | -14.9755
8.5887 | -34.3452 | -14.9853
8.5948 | -34.3455 | -14.9951
8.6009 | -34.3459 | -15.0048
8.607 | -34.3463 | -15.0147
8.6131 | -34.3467 | -15.0245
8.6192 | -34.3471 | -15.0342
8.6253 | -34.3475 | -15.0439
8.6314 | -34.3478 | -15.0536
8.6375 | -34.3482 | -15.0635
8.6436 | -34.3486 | -15.0733
8.6497 | -34.3489 | -15.083
8.6558 | -34.3493 | -15.0928
8.6619 | -34.3497 | -15.1026
8.6681 | -34.35 | -15.1126
8.6742 | -34.3504 | -15.1223
8.6803 | -34.3508 | -15.1321
8.6864 | -34.3512 | -15.1416
8.6925 | -34.3516 | -15.1512
8.6986 | -34.352 | -15.161
8.7047 | -34.3523 | -15.171
8.7108 | -34.3527 | -15.1807
8.7169 | -34.3531 | -15.1905
8.723 | -34.3534 | -15.2002
8.7291 | -34.3538 | -15.2098
8.7352 | -34.3542 | -15.2197
8.7413 | -34.3546 | -15.2296
8.7474 | -34.3549 | -15.2392
8.7535 | -34.3553 | -15.249
8.7596 | -34.3557 | -15.2588
8.7657 | -34.3561 | -15.2686
8.7718 | -34.3564 | -15.2785
8.7779 | -34.3568 | -15.2883
8.784 | -34.3572 | -15.2981
8.7901 | -34.3576 | -15.3079
8.7962 | -34.358 | -15.3174
8.8023 | -34.3584 | -15.3271
8.8084 | -34.3587 | -15.3368
8.8146 | -34.3591 | -15.3466
8.8207 | -34.3595 | -15.3563
8.8268 | -34.3599 | -15.3657
8.8329 | -34.3603 | -15.3755
8.839 | -34.3607 | -15.3852
8.8451 | -34.361 | -15.3949
8.8512 | -34.3614 | -15.4045
8.8573 | -34.3618 | -15.4142
8.8634 | -34.3622 | -15.4238
8.8695 | -34.3626 | -15.4337
8.8756 | -34.3629 | -15.4435
8.8817 | -34.3633 | -15.4531
8.8878 | -34.3637 | -15.463
8.8939 | -34.3641 | -15.4727
8.9 | -34.3645 | -15.4824
8.9061 | -34.3648 | -15.4921
8.9122 | -34.3652 | -15.5018
8.9183 | -34.3656 | -15.5116
8.9244 | -34.366 | -15.5215
8.9305 | -34.3665 | -15.5309
8.9366 | -34.3669 | -15.5406
8.9427 | -34.3672 | -15.5504
8.9488 | -34.3676 | -15.5601
8.955 | -34.368 | -15.5697
8.9611 | -34.3684 | -15.5796
8.9672 | -34.3688 | -15.5893
8.9733 | -34.3692 | -15.5988
8.9794 | -34.3696 | -15.6084
8.9855 | -34.37 | -15.6182
8.9916 | -34.3704 | -15.6281
8.9977 | -34.3708 | -15.6377
9.0038 | -34.3712 | -15.6473
9.0099 | -34.3716 | -15.657
9.016 | -34.372 | -15.6667
9.0221 | -34.3724 | -15.6762
9.0282 | -34.3728 | -15.6857
9.0343 | -34.3732 | -15.6954
9.0404 | -34.3736 | -15.7051
9.0465 | -34.374 | -15.7149
9.0526 | -34.3744 | -15.7245
9.0587 | -34.3748 | -15.7342
9.0648 | -34.3752 | -15.7437
9.0709 | -34.3756 | -15.7534
9.077 | -34.376 | -15.7631
9.0831 | -34.3764 | -15.7729
9.0892 | -34.3768 | -15.7826
9.0953 | -34.3772 | -15.7924
9.1015 | -34.3776 | -15.802
9.1076 | -34.378 | -15.8116
9.1137 | -34.3784 | -15.8212
9.1198 | -34.3788 | -15.8308
9.1259 | -34.3792 | -15.8404
9.132 | -34.3796 | -15.8503
9.1381 | -34.3801 | -15.8598
9.1442 | -34.3805 | -15.8692
9.1503 | -34.3809 | -15.879
9.1564 | -34.3813 | -15.8886
9.1625 | -34.3817 | -15.8983
9.1686 | -34.3821 | -15.9079
9.1747 | -34.3825 | -15.9175
9.1808 | -34.3829 | -15.9272
9.1869 | -34.3834 | -15.9368
9.193 | -34.3837 | -15.9465
9.1991 | -34.3841 | -15.9562
9.2052 | -34.3846 | -15.9658
9.2113 | -34.385 | -15.9755
9.2174 | -34.3854 | -15.9852
9.2235 | -34.3858 | -15.9948
9.2296 | -34.3863 | -16.0044
9.2357 | -34.3867 | -16.0142
9.2419 | -34.3871 | -16.0238
9.248 | -34.3875 | -16.0334
9.2541 | -34.3879 | -16.0431
9.2602 | -34.3884 | -16.0529
9.2663 | -34.3888 | -16.0623
9.2724 | -34.3892 | -16.0719
9.2785 | -34.3896 | -16.0815
9.2846 | -34.39 | -16.091
9.2907 | -34.3904 | -16.1006
9.2968 | -34.3908 | -16.1104
9.3029 | -34.3913 | -16.1201
9.309 | -34.3917 | -16.1297
9.3151 | -34.3921 | -16.1392
9.3212 | -34.3925 | -16.1487
9.3273 | -34.3929 | -16.1583
9.3334 | -34.3934 | -16.1678
9.3395 | -34.3938 | -16.1773
9.3456 | -34.3942 | -16.1868
9.3517 | -34.3946 | -16.1965
9.3578 | -34.395 | -16.2061
9.3639 | -34.3954 | -16.2155
9.37 | -34.3958 | -16.2251
9.3761 | -34.3963 | -16.2349
9.3822 | -34.3967 | -16.2446
9.3884 | -34.3971 | -16.254
9.3945 | -34.3976 | -16.2636
9.4006 | -34.398 | -16.2732
9.4067 | -34.3984 | -16.2828
9.4128 | -34.3988 | -16.2925
9.4189 | -34.3993 | -16.3023
9.425 | -34.3997 | -16.3118
9.4311 | -34.4001 | -16.3214
9.4372 | -34.4006 | -16.3309
9.4433 | -34.401 | -16.3405
9.4494 | -34.4014 | -16.3499
9.4555 | -34.4019 | -16.3596
9.4616 | -34.4023 | -16.3692
9.4677 | -34.4027 | -16.3788
9.4738 | -34.4031 | -16.3884
9.4799 | -34.4036 | -16.398
9.486 | -34.404 | -16.4076
9.4921 | -34.4044 | -16.4171
9.4982 | -34.4049 | -16.4265
9.5043 | -34.4053 | -16.4361
9.5104 | -34.4057 | -16.4455
9.5165 | -34.4061 | -16.4551
9.5226 | -34.4065 | -16.4649
9.5288 | -34.407 | -16.4743
9.5349 | -34.4074 | -16.4839
9.541 | -34.4078 | -16.4933
9.5471 | -34.4082 | -16.5028
9.5532 | -34.4087 | -16.5125
9.5593 | -34.4091 | -16.5222
9.5654 | -34.4096 | -16.5317
9.5715 | -34.41 | -16.5412
9.5776 | -34.4104 | -16.5509
9.5837 | -34.4108 | -16.5604
9.5898 | -34.4112 | -16.5701
9.5959 | -34.4116 | -16.5797
9.602 | -34.4121 | -16.589
9.6081 | -34.4125 | -16.5986
9.6142 | -34.4129 | -16.608
9.6203 | -34.4133 | -16.6176
9.6264 | -34.4138 | -16.6271
9.6325 | -34.4142 | -16.6367
9.6386 | -34.4146 | -16.6462
9.6447 | -34.415 | -16.6558
9.6508 | -34.4154 | -16.6653
9.6569 | -34.4159 | -16.6748
9.663 | -34.4163 | -16.6845
9.6691 | -34.4168 | -16.694
9.6753 | -34.4172 | -16.7035
9.6814 | -34.4176 | -16.7129
9.6875 | -34.4181 | -16.7222
9.6936 | -34.4185 | -16.7316
9.6997 | -34.4189 | -16.7412
9.7058 | -34.4194 | -16.7507
9.7119 | -34.4198 | -16.7604
9.718 | -34.4203 | -16.77
9.7241 | -34.4207 | -16.7794
9.7302 | -34.4211 | -16.789
9.7363 | -34.4216 | -16.7984
9.7424 | -34.422 | -16.8077
9.7485 | -34.4224 | -16.8173
9.7546 | -34.4228 | -16.8267
9.7607 | -34.4232 | -16.8364
9.7668 | -34.4237 | -16.846
9.7729 | -34.4241 | -16.8554
9.779 | -34.4246 | -16.8647
9.7851 | -34.425 | -16.8744
9.7912 | -34.4255 | -16.8838
9.7973 | -34.4259 | -16.8934
9.8034 | -34.4264 | -16.9027
9.8095 | -34.4268 | -16.9124
9.8157 | -34.4273 | -16.9219
9.8218 | -34.4277 | -16.9316
9.8279 | -34.4281 | -16.9409
9.834 | -34.4285 | -16.9506
9.8401 | -34.429 | -16.9601
9.8462 | -34.4294 | -16.9693
9.8523 | -34.4298 | -16.9788
9.8584 | -34.4303 | -16.9883
9.8645 | -34.4307 | -16.998
9.8706 | -34.4311 | -17.0075
9.8767 | -34.4315 | -17.017
9.8828 | -34.4319 | -17.0263
9.8889 | -34.4324 | -17.0356
9.895 | -34.4328 | -17.0451
9.9011 | -34.4332 | -17.0546
9.9072 | -34.4337 | -17.0641
9.9133 | -34.4341 | -17.0736
9.9194 | -34.4345 | -17.083
9.9255 | -34.4349 | -17.0924
9.9316 | -34.4354 | -17.1019
9.9377 | -34.4358 | -17.1113
9.9438 | -34.4362 | -17.1208
9.9499 | -34.4367 | -17.1304
9.956 | -34.4371 | -17.1399
9.9622 | -34.4375 | -17.1494
9.9683 | -34.438 | -17.1589
9.9744 | -34.4384 | -17.1686
9.9805 | -34.4388 | -17.178
9.9866 | -34.4392 | -17.1873
9.9927 | -34.4397 | -17.1967
9.9988 | -34.4401 | -17.2063
10.0049 | -34.4405 | -17.2158
10.011 | -34.441 | -17.2253
10.0171 | -34.4414 | -17.2346
10.0232 | -34.4419 | -17.244
10.0293 | -34.4423 | -17.2535
10.0354 | -34.4428 | -17.263
10.0415 | -34.4432 | -17.2724
10.0476 | -34.4436 | -17.282
10.0537 | -34.4441 | -17.2916
10.0598 | -34.4445 | -17.3009
10.0659 | -34.4449 | -17.3104
10.072 | -34.4453 | -17.3197
10.0781 | -34.4458 | -17.3291
10.0842 | -34.4462 | -17.3386
10.0903 | -34.4467 | -17.348
10.0964 | -34.4471 | -17.3572
10.1026 | -34.4475 | -17.3667
10.1087 | -34.4479 | -17.3761
10.1148 | -34.4483 | -17.3856
10.1209 | -34.4488 | -17.3949
10.127 | -34.4492 | -17.4042
10.1331 | -34.4497 | -17.4135
10.1392 | -34.4501 | -17.4228
10.1453 | -34.4505 | -17.4321
10.1514 | -34.451 | -17.4416
10.1575 | -34.4514 | -17.4511
10.1636 | -34.4518 | -17.4602
10.1697 | -34.4522 | -17.4698
10.1758 | -34.4527 | -17.479
10.1819 | -34.4531 | -17.4884
10.188 | -34.4535 | -17.4978
10.1941 | -34.454 | -17.5073
10.2002 | -34.4544 | -17.5167
10.2063 | -34.4548 | -17.526
10.2124 | -34.4552 | -17.5355
10.2185 | -34.4557 | -17.545
10.2246 | -34.4561 | -17.5542
10.2307 | -34.4565 | -17.5637
10.2368 | -34.457 | -17.5728
10.2429 | -34.4574 | -17.5822
10.2491 | -34.4578 | -17.5916
10.2552 | -34.4582 | -17.601
10.2613 | -34.4587 | -17.6105
10.2674 | -34.4591 | -17.6199
10.2735 | -34.4596 | -17.6294
10.2796 | -34.46 | -17.6388
10.2857 | -34.4604 | -17.6483
10.2918 | -34.4609 | -17.6578
10.2979 | -34.4613 | -17.6672
10.304 | -34.4618 | -17.6763
10.3101 | -34.4622 | -17.6858
10.3162 | -34.4626 | -17.6953
10.3223 | -34.463 | -17.7045
10.3284 | -34.4635 | -17.7139
10.3345 | -34.4639 | -17.7231
10.3406 | -34.4643 | -17.7324
10.3467 | -34.4647 | -17.7417
10.3528 | -34.4652 | -17.751
10.3589 | -34.4656 | -17.7604
10.365 | -34.466 | -17.7699
10.3711 | -34.4665 | -17.7793
10.3772 | -34.4669 | -17.7887
10.3833 | -34.4674 | -17.798
10.3895 | -34.4678 | -17.8073
10.3956 | -34.4682 | -17.8167
10.4017 | -34.4687 | -17.8259
10.4078 | -34.4691 | -17.8353
10.4139 | -34.4695 | -17.8446
10.42 | -34.47 | -17.8541
10.4261 | -34.4704 | -17.8634
10.4322 | -34.4708 | -17.8728
10.4383 | -34.4712 | -17.8821
10.4444 | -34.4717 | -17.8913
10.4505 | -34.4721 | -17.9007
10.4566 | -34.4725 | -17.9099
10.4627 | -34.4729 | -17.9191
10.4688 | -34.4734 | -17.9283
10.4749 | -34.4738 | -17.9378
10.481 | -34.4742 | -17.947
10.4871 | -34.4747 | -17.9563
10.4932 | -34.4751 | -17.9655
10.4993 | -34.4755 | -17.9748
10.5054 | -34.476 | -17.9841
10.5115 | -34.4764 | -17.9935
10.5176 | -34.4769 | -18.0026
10.5237 | -34.4773 | -18.0121
10.5298 | -34.4777 | -18.0213
10.536 | -34.4782 | -18.0307
10.5421 | -34.4786 | -18.0399
10.5482 | -34.479 | -18.0493
10.5543 | -34.4795 | -18.0584
10.5604 | -34.4799 | -18.0676
10.5665 | -34.4804 | -18.0768
10.5726 | -34.4808 | -18.086
10.5787 | -34.4812 | -18.0953
10.5848 | -34.4816 | -18.1043
10.5909 | -34.4821 | -18.1137
10.597 | -34.4825 | -18.1231
10.6031 | -34.4829 | -18.1325
10.6092 | -34.4833 | -18.1418
10.6153 | -34.4838 | -18.1512
10.6214 | -34.4842 | -18.1603
10.6275 | -34.4846 | -18.1695
10.6336 | -34.4851 | -18.1788
10.6397 | -34.4855 | -18.1881
10.6458 | -34.4859 | -18.1974
10.6519 | -34.4864 | -18.2067
10.658 | -34.4868 | -18.216
10.6641 | -34.4872 | -18.2254
10.6702 | -34.4877 | -18.2347
10.6764 | -34.4881 | -18.2441
10.6825 | -34.4885 | -18.2532
10.6886 | -34.489 | -18.2623
10.6947 | -34.4894 | -18.2715
10.7008 | -34.4899 | -18.2808
10.7069 | -34.4903 | -18.2902
10.713 | -34.4907 | -18.2994
10.7191 | -34.4912 | -18.3087
10.7252 | -34.4916 | -18.3175
10.7313 | -34.492 | -18.3267
10.7374 | -34.4924 | -18.336
10.7435 | -34.4929 | -18.3453
10.7496 | -34.4933 | -18.3543
10.7557 | -34.4937 | -18.3637
10.7618 | -34.4942 | -18.3727
10.7679 | -34.4946 | -18.3819
10.774 | -34.495 | -18.3913
10.7801 | -34.4955 | -18.4005
10.7862 | -34.4959 | -18.4096
10.7923 | -34.4963 | -18.4188
10.7984 | -34.4968 | -18.4279
10.8045 | -34.4972 | -18.4369
10.8106 | -34.4976 | -18.4461
10.8168 | -34.498 | -18.4552
10.8229 | -34.4985 | -18.4644
10.829 | -34.4989 | -18.4736
10.8351 | -34.4993 | -18.4825
10.8412 | -34.4998 | -18.4916
10.8473 | -34.5002 | -18.5006
10.8534 | -34.5007 | -18.5098
10.8595 | -34.5011 | -18.519
10.8656 | -34.5016 | -18.5279
10.8717 | -34.502 | -18.5371
10.8778 | -34.5024 | -18.5463
10.8839 | -34.5029 | -18.5555
10.89 | -34.5033 | -18.5647
10.8961 | -34.5038 | -18.5739
10.9022 | -34.5042 | -18.5829
10.9083 | -34.5046 | -18.5921
10.9144 | -34.5051 | -18.6011
10.9205 | -34.5055 | -18.6103
10.9266 | -34.5059 | -18.6194
10.9327 | -34.5064 | -18.6285
10.9388 | -34.5068 | -18.6375
10.9449 | -34.5073 | -18.6466
10.951 | -34.5077 | -18.6558
10.9571 | -34.5081 | -18.665
10.9633 | -34.5086 | -18.6742
10.9694 | -34.509 | -18.6835
10.9755 | -34.5094 | -18.6927
10.9816 | -34.5099 | -18.702
10.9877 | -34.5103 | -18.711
10.9938 | -34.5107 | -18.7202
10.9999 | -34.5112 | -18.7292
11.006 | -34.5116 | -18.7382
11.0121 | -34.5121 | -18.7473
11.0182 | -34.5125 | -18.7564
11.0243 | -34.513 | -18.7655
11.0304 | -34.5134 | -18.7747
11.0365 | -34.5138 | -18.7836
11.0426 | -34.5143 | -18.7926
11.0487 | -34.5148 | -18.8018
11.0548 | -34.5152 | -18.8108
11.0609 | -34.5157 | -18.82
11.067 | -34.5161 | -18.8291
11.0731 | -34.5165 | -18.8381
11.0792 | -34.517 | -18.8471
11.0853 | -34.5174 | -18.8563
11.0914 | -34.5179 | -18.8655
11.0975 | -34.5183 | -18.8745
11.1037 | -34.5187 | -18.8835
11.1098 | -34.5192 | -18.8928
11.1159 | -34.5196 | -18.9016
11.122 | -34.5201 | -18.9107
11.1281 | -34.5205 | -18.9198
11.1342 | -34.521 | -18.9289
11.1403 | -34.5214 | -18.9381
11.1464 | -34.5218 | -18.9472
11.1525 | -34.5223 | -18.9561
11.1586 | -34.5227 | -18.9653
11.1647 | -34.5232 | -18.9742
11.1708 | -34.5236 | -18.9833
11.1769 | -34.524 | -18.9924
11.183 | -34.5245 | -19.0014
11.1891 | -34.5249 | -19.0104
11.1952 | -34.5254 | -19.0195
11.2013 | -34.5259 | -19.0285
11.2074 | -34.5263 | -19.0376
11.2135 | -34.5267 | -19.0468
11.2196 | -34.5272 | -19.0558
11.2257 | -34.5276 | -19.0649
11.2318 | -34.5281 | -19.0738
11.2379 | -34.5286 | -19.083
11.244 | -34.529 | -19.0921
11.2502 | -34.5295 | -19.1011
11.2563 | -34.53 | -19.1102
11.2624 | -34.5304 | -19.1191
11.2685 | -34.5308 | -19.1281
11.2746 | -34.5313 | -19.137
11.2807 | -34.5317 | -19.1462
11.2868 | -34.5322 | -19.1554
11.2929 | -34.5326 | -19.1641
11.299 | -34.5331 | -19.1732
11.3051 | -34.5335 | -19.1825
11.3112 | -34.534 | -19.1915
11.3173 | -34.5344 | -19.2004
11.3234 | -34.5349 | -19.2092
11.3295 | -34.5353 | -19.2183
11.3356 | -34.5358 | -19.2275
11.3417 | -34.5362 | -19.2364
11.3478 | -34.5366 | -19.2455
11.3539 | -34.5371 | -19.2547
11.36 | -34.5375 | -19.2635
11.3661 | -34.538 | -19.2724
11.3722 | -34.5384 | -19.2813
11.3783 | -34.5389 | -19.2904
11.3844 | -34.5393 | -19.2994
11.3906 | -34.5398 | -19.3084
11.3967 | -34.5402 | -19.3172
11.4028 | -34.5407 | -19.3264
11.4089 | -34.5411 | -19.3354
11.415 | -34.5415 | -19.3444
11.4211 | -34.542 | -19.3534
11.4272 | -34.5424 | -19.3624
11.4333 | -34.5428 | -19.3713
11.4394 | -34.5433 | -19.3803
11.4455 | -34.5437 | -19.3892
11.4516 | -34.5441 | -19.3983
11.4577 | -34.5446 | -19.4073
11.4638 | -34.5451 | -19.4159
11.4699 | -34.5455 | -19.425
11.476 | -34.5459 | -19.4341
11.4821 | -34.5464 | -19.443
11.4882 | -34.5469 | -19.4521
11.4943 | -34.5473 | -19.4611
11.5004 | -34.5478 | -19.4698
11.5065 | -34.5482 | -19.4788
11.5126 | -34.5487 | -19.4876
11.5187 | -34.5491 | -19.4965
11.5248 | -34.5495 | -19.5057
11.5309 | -34.55 | -19.5146
11.5371 | -34.5505 | -19.5235
11.5432 | -34.5509 | -19.5325
11.5493 | -34.5513 | -19.5417
11.5554 | -34.5518 | -19.5506
11.5615 | -34.5522 | -19.5597
11.5676 | -34.5527 | -19.5687
11.5737 | -34.5531 | -19.5773
11.5798 | -34.5536 | -19.5862
11.5859 | -34.554 | -19.5951
11.592 | -34.5545 | -19.6039
11.5981 | -34.5549 | -19.6129
11.6042 | -34.5554 | -19.6217
11.6103 | -34.5558 | -19.6306
11.6164 | -34.5563 | -19.6393
11.6225 | -34.5567 | -19.6482
11.6286 | -34.5571 | -19.6571
11.6347 | -34.5576 | -19.6661
11.6408 | -34.558 | -19.675
11.6469 | -34.5585 | -19.684
11.653 | -34.5589 | -19.6928
11.6591 | -34.5594 | -19.7017
11.6652 | -34.5598 | -19.7105
11.6713 | -34.5603 | -19.7193
11.6775 | -34.5607 | -19.7281
11.6836 | -34.5612 | -19.737
11.6897 | -34.5616 | -19.7459
11.6958 | -34.5621 | -19.7548
11.7019 | -34.5625 | -19.7639
11.708 | -34.563 | -19.7727
11.7141 | -34.5634 | -19.7816
11.7202 | -34.5639 | -19.7904
11.7263 | -34.5643 | -19.7994
11.7324 | -34.5647 | -19.8084
11.7385 | -34.5652 | -19.8174
11.7446 | -34.5656 | -19.8261
11.7507 | -34.5661 | -19.8348
11.7568 | -34.5665 | -19.8437
11.7629 | -34.567 | -19.8526
11.769 | -34.5674 | -19.8614
11.7751 | -34.5678 | -19.8702
11.7812 | -34.5683 | -19.8789
11.7873 | -34.5688 | -19.8877
11.7934 | -34.5692 | -19.8966
11.7995 | -34.5697 | -19.9054
11.8056 | -34.5702 | -19.9142
11.8117 | -34.5706 | -19.923
11.8178 | -34.5711 | -19.9318
11.824 | -34.5715 | -19.9405
11.8301 | -34.572 | -19.9493
11.8362 | -34.5724 | -19.9581
11.8423 | -34.5729 | -19.967
11.8484 | -34.5733 | -19.9758
11.8545 | -34.5738 | -19.9844
11.8606 | -34.5743 | -19.9932
11.8667 | -34.5747 | -20.0021
11.8728 | -34.5752 | -20.0107
11.8789 | -34.5756 | -20.0196
11.885 | -34.5761 | -20.0284
11.8911 | -34.5765 | -20.0372
11.8972 | -34.577 | -20.046
11.9033 | -34.5774 | -20.0548
11.9094 | -34.5779 | -20.0636
11.9155 | -34.5783 | -20.0724
11.9216 | -34.5788 | -20.0811
11.9277 | -34.5793 | -20.0899
11.9338 | -34.5797 | -20.0989
11.9399 | -34.5802 | -20.1079
11.946 | -34.5806 | -20.1166
11.9521 | -34.581 | -20.1253
11.9582 | -34.5815 | -20.134
11.9644 | -34.5819 | -20.1426
11.9705 | -34.5824 | -20.1515
11.9766 | -34.5829 | -20.16
11.9827 | -34.5833 | -20.1688
11.9888 | -34.5838 | -20.1776
11.9949 | -34.5843 | -20.186
12.001 | -34.5847 | -20.1948
12.0071 | -34.5852 | -20.2035
12.0132 | -34.5856 | -20.2122
12.0193 | -34.5861 | -20.2211
12.0254 | -34.5865 | -20.2298
12.0315 | -34.587 | -20.2381
12.0376 | -34.5874 | -20.2469
12.0437 | -34.5879 | -20.2554
12.0498 | -34.5884 | -20.2642
12.0559 | -34.5888 | -20.2729
12.062 | -34.5893 | -20.2814
12.0681 | -34.5898 | -20.29
12.0742 | -34.5902 | -20.2987
12.0803 | -34.5907 | -20.3076
12.0864 | -34.5911 | -20.3165
12.0925 | -34.5916 | -20.3252
12.0986 | -34.592 | -20.334
12.1047 | -34.5925 | -20.3427
12.1109 | -34.5929 | -20.3511
12.117 | -34.5934 | -20.3599
12.1231 | -34.5939 | -20.3686
12.1292 | -34.5943 | -20.3772
12.1353 | -34.5948 | -20.3857
12.1414 | -34.5952 | -20.3944
12.1475 | -34.5957 | -20.4032
12.1536 | -34.5961 | -20.4118
12.1597 | -34.5966 | -20.4206
12.1658 | -34.5971 | -20.4293
12.1719 | -34.5975 | -20.4378
12.178 | -34.598 | -20.4465
12.1841 | -34.5985 | -20.4552
12.1902 | -34.5989 | -20.4638
12.1963 | -34.5994 | -20.4725
12.2024 | -34.5998 | -20.481
12.2085 | -34.6003 | -20.4897
12.2146 | -34.6007 | -20.4984
12.2207 | -34.6012 | -20.5072
12.2268 | -34.6017 | -20.5157
12.2329 | -34.6021 | -20.5244
12.239 | -34.6026 | -20.5326
12.2451 | -34.6031 | -20.5414
12.2513 | -34.6035 | -20.55
12.2574 | -34.604 | -20.5588
12.2635 | -34.6045 | -20.5675
12.2696 | -34.6049 | -20.5762
12.2757 | -34.6054 | -20.5845
12.2818 | -34.6059 | -20.5932
12.2879 | -34.6063 | -20.6018
12.294 | -34.6068 | -20.6103
12.3001 | -34.6072 | -20.619
12.3062 | -34.6077 | -20.6276
12.3123 | -34.6081 | -20.636
12.3184 | -34.6086 | -20.6444
12.3245 | -34.609 | -20.653
12.3306 | -34.6095 | -20.6615
12.3367 | -34.6099 | -20.6703
12.3428 | -34.6104 | -20.6787
12.3489 | -34.6109 | -20.6869
12.355 | -34.6113 | -20.6957
12.3611 | -34.6118 | -20.7044
12.3672 | -34.6122 | -20.713
12.3733 | -34.6127 | -20.7216
12.3794 | -34.6132 | -20.7299
12.3855 | -34.6136 | -20.7383
12.3916 | -34.6141 | -20.7469
12.3978 | -34.6146 | -20.7555
12.4039 | -34.6151 | -20.764
12.41 | -34.6155 | -20.7727
12.4161 | -34.616 | -20.7809
12.4222 | -34.6165 | -20.7895
12.4283 | -34.6169 | -20.798
12.4344 | -34.6174 | -20.8067
12.4405 | -34.6179 | -20.8153
12.4466 | -34.6184 | -20.8239
12.4527 | -34.6188 | -20.8323
12.4588 | -34.6193 | -20.8409
12.4649 | -34.6197 | -20.8495
12.471 | -34.6202 | -20.8581
12.4771 | -34.6207 | -20.8666
12.4832 | -34.6211 | -20.875
12.4893 | -34.6216 | -20.8833
12.4954 | -34.622 | -20.8917
12.5015 | -34.6225 | -20.9003
12.5076 | -34.6229 | -20.909
12.5137 | -34.6234 | -20.9175
12.5198 | -34.6239 | -20.926
12.5259 | -34.6244 | -20.9347
12.532 | -34.6248 | -20.9432
12.5382 | -34.6253 | -20.9518
12.5443 | -34.6258 | -20.9602
12.5504 | -34.6262 | -20.9687
12.5565 | -34.6267 | -20.977
12.5626 | -34.6272 | -20.9857
12.5687 | -34.6276 | -20.9942
12.5748 | -34.6281 | -21.0028
12.5809 | -34.6285 | -21.0114
12.587 | -34.629 | -21.02
12.5931 | -34.6295 | -21.0282
12.5992 | -34.63 | -21.0368
12.6053 | -34.6304 | -21.0452
12.6114 | -34.6309 | -21.0537
12.6175 | -34.6314 | -21.0623
12.6236 | -34.6318 | -21.0707
12.6297 | -34.6323 | -21.0791
12.6358 | -34.6328 | -21.0874
12.6419 | -34.6332 | -21.0959
12.648 | -34.6337 | -21.1045
12.6541 | -34.6342 | -21.113
12.6602 | -34.6346 | -21.1213
12.6663 | -34.6351 | -21.1299
12.6724 | -34.6356 | -21.1383
12.6785 | -34.636 | -21.1468
12.6847 | -34.6365 | -21.155
12.6908 | -34.637 | -21.1636
12.6969 | -34.6374 | -21.1719
12.703 | -34.6379 | -21.1805
12.7091 | -34.6384 | -21.189
12.7152 | -34.6388 | -21.1974
12.7213 | -34.6393 | -21.2059
12.7274 | -34.6398 | -21.2145
12.7335 | -34.6403 | -21.2228
12.7396 | -34.6407 | -21.2311
12.7457 | -34.6412 | -21.2397
12.7518 | -34.6416 | -21.2482
12.7579 | -34.6421 | -21.2568
12.764 | -34.6426 | -21.2653
12.7701 | -34.6431 | -21.2737
12.7762 | -34.6435 | -21.2821
12.7823 | -34.644 | -21.2906
12.7884 | -34.6445 | -21.2991
12.7945 | -34.6449 | -21.3077
12.8006 | -34.6454 | -21.3161
12.8067 | -34.6459 | -21.3246
12.8128 | -34.6463 | -21.333
12.8189 | -34.6468 | -21.3415
12.8251 | -34.6473 | -21.3499
12.8312 | -34.6477 | -21.3585
12.8373 | -34.6482 | -21.367
12.8434 | -34.6487 | -21.3755
12.8495 | -34.6491 | -21.384
12.8556 | -34.6496 | -21.3923
12.8617 | -34.6501 | -21.4007
12.8678 | -34.6506 | -21.4093
12.8739 | -34.651 | -21.4177
12.88 | -34.6515 | -21.4261
12.8861 | -34.6519 | -21.4345
12.8922 | -34.6524 | -21.4431
12.8983 | -34.6529 | -21.4515
12.9044 | -34.6533 | -21.4599
12.9105 | -34.6538 | -21.4682
12.9166 | -34.6543 | -21.4767
12.9227 | -34.6548 | -21.4852
12.9288 | -34.6552 | -21.4936
12.9349 | -34.6557 | -21.502
12.941 | -34.6562 | -21.5104
12.9471 | -34.6567 | -21.5188
12.9532 | -34.6572 | -21.5272
12.9593 | -34.6576 | -21.5358
12.9654 | -34.6581 | -21.5444
12.9716 | -34.6586 | -21.553
12.9777 | -34.6591 | -21.5614
12.9838 | -34.6596 | -21.5697
12.9899 | -34.66 | -21.5782
12.996 | -34.6605 | -21.5866
13.0021 | -34.661 | -21.595
13.0082 | -34.6615 | -21.6034
13.0143 | -34.662 | -21.6116
13.0204 | -34.6624 | -21.6199
13.0265 | -34.6629 | -21.6283
13.0326 | -34.6634 | -21.6368
13.0387 | -34.6639 | -21.6453
13.0448 | -34.6643 | -21.6537
13.0509 | -34.6648 | -21.662
13.057 | -34.6653 | -21.6704
13.0631 | -34.6658 | -21.6787
13.0692 | -34.6663 | -21.6871
13.0753 | -34.6667 | -21.6956
13.0814 | -34.6672 | -21.7041
13.0875 | -34.6677 | -21.7125
13.0936 | -34.6682 | -21.721
13.0997 | -34.6686 | -21.7293
13.1058 | -34.6691 | -21.7377
13.112 | -34.6696 | -21.746
13.1181 | -34.6701 | -21.7543
13.1242 | -34.6705 | -21.7627
13.1303 | -34.671 | -21.7711
13.1364 | -34.6715 | -21.7795
13.1425 | -34.6719 | -21.7878
13.1486 | -34.6724 | -21.7962
13.1547 | -34.6729 | -21.8046
13.1608 | -34.6733 | -21.8129
13.1669 | -34.6738 | -21.8212
13.173 | -34.6743 | -21.8297
13.1791 | -34.6748 | -21.838
13.1852 | -34.6752 | -21.8465
13.1913 | -34.6757 | -21.8549
13.1974 | -34.6762 | -21.8634
13.2035 | -34.6766 | -21.8718
13.2096 | -34.6771 | -21.8801
13.2157 | -34.6776 | -21.8883
13.2218 | -34.6781 | -21.8967
13.2279 | -34.6786 | -21.9051
13.234 | -34.679 | -21.9134
13.2401 | -34.6795 | -21.9218
13.2462 | -34.68 | -21.9302
13.2524 | -34.6804 | -21.9388
13.2585 | -34.6809 | -21.947
13.2646 | -34.6814 | -21.9555
13.2707 | -34.6819 | -21.9638
13.2768 | -34.6824 | -21.9723
13.2829 | -34.6829 | -21.9808
13.289 | -34.6833 | -21.9891
13.2951 | -34.6838 | -21.9975
13.3012 | -34.6843 | -22.0059
13.3073 | -34.6848 | -22.0144
13.3134 | -34.6852 | -22.0229
13.3195 | -34.6857 | -22.0312
13.3256 | -34.6862 | -22.0395
13.3317 | -34.6867 | -22.0477
13.3378 | -34.6872 | -22.0562
13.3439 | -34.6876 | -22.0646
13.35 | -34.6881 | -22.0731
13.3561 | -34.6886 | -22.0814
13.3622 | -34.689 | -22.0897
13.3683 | -34.6895 | -22.098
13.3744 | -34.69 | -22.1063
13.3805 | -34.6905 | -22.1148
13.3866 | -34.6909 | -22.123
13.3927 | -34.6914 | -22.1314
13.3989 | -34.6919 | -22.1398
13.405 | -34.6923 | -22.1481
13.4111 | -34.6928 | -22.1564
13.4172 | -34.6933 | -22.1648
13.4233 | -34.6938 | -22.1733
13.4294 | -34.6942 | -22.1819
13.4355 | -34.6947 | -22.1901
13.4416 | -34.6952 | -22.1984
13.4477 | -34.6957 | -22.2068
13.4538 | -34.6962 | -22.2151
13.4599 | -34.6967 | -22.2235
13.466 | -34.6971 | -22.2319
13.4721 | -34.6976 | -22.2403
13.4782 | -34.6981 | -22.2487
13.4843 | -34.6986 | -22.2571
13.4904 | -34.6991 | -22.2654
13.4965 | -34.6995 | -22.2739
13.5026 | -34.7 | -22.2823
13.5087 | -34.7004 | -22.2905
13.5148 | -34.7009 | -22.299
13.5209 | -34.7014 | -22.3075
13.527 | -34.7019 | -22.3159
13.5331 | -34.7023 | -22.3242
13.5393 | -34.7029 | -22.3325
13.5454 | -34.7033 | -22.3408
13.5515 | -34.7038 | -22.3493
13.5576 | -34.7043 | -22.3577
13.5637 | -34.7048 | -22.3661
13.5698 | -34.7053 | -22.3744
13.5759 | -34.7057 | -22.3826
13.582 | -34.7062 | -22.3909
13.5881 | -34.7067 | -22.3993
13.5942 | -34.7072 | -22.4078
13.6003 | -34.7077 | -22.4162
13.6064 | -34.7081 | -22.4247
13.6125 | -34.7086 | -22.4331
13.6186 | -34.7091 | -22.4413
13.6247 | -34.7096 | -22.4498
13.6308 | -34.7101 | -22.4582
13.6369 | -34.7106 | -22.4666
13.643 | -34.711 | -22.4751
13.6491 | -34.7115 | -22.4834
13.6552 | -34.712 | -22.4917
13.6613 | -34.7125 | -22.5001
13.6674 | -34.713 | -22.5085
13.6735 | -34.7134 | -22.5169
13.6796 | -34.7139 | -22.5252
13.6858 | -34.7144 | -22.5336
13.6919 | -34.7149 | -22.542
13.698 | -34.7154 | -22.5504
13.7041 | -34.7159 | -22.5589
13.7102 | -34.7164 | -22.5672
13.7163 | -34.7169 | -22.5756
13.7224 | -34.7173 | -22.5839
13.7285 | -34.7178 | -22.5923
13.7346 | -34.7183 | -22.6008
13.7407 | -34.7188 | -22.6091
13.7468 | -34.7193 | -22.6175
13.7529 | -34.7198 | -22.6257
13.759 | -34.7202 | -22.6342
13.7651 | -34.7207 | -22.6427
13.7712 | -34.7212 | -22.6511
13.7773 | -34.7217 | -22.6595
13.7834 | -34.7221 | -22.6678
13.7895 | -34.7227 | -22.6761
13.7956 | -34.7231 | -22.6846
13.8017 | -34.7236 | -22.6929
13.8078 | -34.7241 | -22.7013
13.8139 | -34.7245 | -22.7095
13.82 | -34.725 | -22.7178
13.8262 | -34.7255 | -22.7262
13.8323 | -34.726 | -22.7347
13.8384 | -34.7265 | -22.7431
13.8445 | -34.7269 | -22.7515
13.8506 | -34.7274 | -22.76
13.8567 | -34.7279 | -22.7685
13.8628 | -34.7284 | -22.7769
13.8689 | -34.7289 | -22.7854
13.875 | -34.7294 | -22.7937
13.8811 | -34.7298 | -22.8021
13.8872 | -34.7303 | -22.8104
13.8933 | -34.7308 | -22.8186
13.8994 | -34.7313 | -22.827
13.9055 | -34.7318 | -22.8355
13.9116 | -34.7322 | -22.8438
13.9177 | -34.7327 | -22.8522
13.9238 | -34.7332 | -22.8606
13.9299 | -34.7337 | -22.8689
13.936 | -34.7341 | -22.8773
13.9421 | -34.7346 | -22.8856
13.9482 | -34.7351 | -22.8939
13.9543 | -34.7356 | -22.9024
13.9604 | -34.736 | -22.9107
13.9665 | -34.7365 | -22.919
13.9727 | -34.737 | -22.9276
13.9788 | -34.7375 | -22.9359
13.9849 | -34.738 | -22.9443
13.991 | -34.7385 | -22.9527
13.9971 | -34.739 | -22.9609
14.0032 | -34.7394 | -22.9693
14.0093 | -34.7399 | -22.9776
14.0154 | -34.7404 | -22.9859
14.0215 | -34.7409 | -22.9942
14.0276 | -34.7414 | -23.0026
14.0337 | -34.7419 | -23.0109
14.0398 | -34.7424 | -23.0192
14.0459 | -34.7429 | -23.0275
14.052 | -34.7433 | -23.036
14.0581 | -34.7438 | -23.0444
14.0642 | -34.7443 | -23.0527
14.0703 | -34.7448 | -23.0611
14.0764 | -34.7453 | -23.0693
14.0825 | -34.7458 | -23.0777
14.0886 | -34.7462 | -23.0861
14.0947 | -34.7467 | -23.0945
14.1008 | -34.7472 | -23.1028
14.1069 | -34.7477 | -23.1111
14.1131 | -34.7482 | -23.1195
14.1192 | -34.7487 | -23.1278
14.1253 | -34.7491 | -23.1361
14.1314 | -34.7496 | -23.1445
14.1375 | -34.7501 | -23.1528
14.1436 | -34.7506 | -23.1611
14.1497 | -34.7511 | -23.1694
14.1558 | -34.7516 | -23.1778
14.1619 | -34.7521 | -23.1861
14.168 | -34.7526 | -23.1944
14.1741 | -34.7531 | -23.2026
14.1802 | -34.7535 | -23.211
14.1863 | -34.754 | -23.2194
14.1924 | -34.7545 | -23.2278
14.1985 | -34.755 | -23.2361
14.2046 | -34.7555 | -23.2446
14.2107 | -34.756 | -23.2527
14.2168 | -34.7565 | -23.2611
14.2229 | -34.757 | -23.2695
14.229 | -34.7575 | -23.278
14.2351 | -34.7579 | -23.2863
14.2412 | -34.7584 | -23.2946
14.2473 | -34.7589 | -23.3028
14.2534 | -34.7594 | -23.3111
14.2596 | -34.7599 | -23.3193
14.2657 | -34.7603 | -23.3276
14.2718 | -34.7608 | -23.336
14.2779 | -34.7613 | -23.3444
14.284 | -34.7618 | -23.3528
14.2901 | -34.7623 | -23.3611
14.2962 | -34.7628 | -23.3694
14.3023 | -34.7633 | -23.3778
14.3084 | -34.7638 | -23.3862
14.3145 | -34.7643 | -23.3945
14.3206 | -34.7648 | -23.403
14.3267 | -34.7652 | -23.4114
14.3328 | -34.7657 | -23.4197
14.3389 | -34.7662 | -23.4279
14.345 | -34.7667 | -23.4363
14.3511 | -34.7672 | -23.4446
14.3572 | -34.7677 | -23.453
14.3633 | -34.7682 | -23.4614
14.3694 | -34.7687 | -23.4697
14.3755 | -34.7692 | -23.478
14.3816 | -34.7697 | -23.4864
14.3877 | -34.7702 | -23.4946
14.3938 | -34.7706 | -23.5031
14.4 | -34.7711 | -23.5113
14.4061 | -34.7716 | -23.5196
14.4122 | -34.7721 | -23.528
14.4183 | -34.7726 | -23.5362
14.4244 | -34.7731 | -23.5446
14.4305 | -34.7736 | -23.5529
14.4366 | -34.7741 | -23.5612
14.4427 | -34.7746 | -23.5696
14.4488 | -34.7751 | -23.578
14.4549 | -34.7756 | -23.5862
14.461 | -34.776 | -23.5944
14.4671 | -34.7765 | -23.6028
14.4732 | -34.777 | -23.6111
14.4793 | -34.7775 | -23.6195
14.4854 | -34.778 | -23.6278
14.4915 | -34.7785 | -23.6358
14.4976 | -34.779 | -23.644
14.5037 | -34.7795 | -23.6522
14.5098 | -34.78 | -23.6605
14.5159 | -34.7804 | -23.6689
14.522 | -34.7809 | -23.6771
14.5281 | -34.7814 | -23.6853
14.5342 | -34.7819 | -23.6937
14.5403 | -34.7824 | -23.7019
14.5465 | -34.7829 | -23.7102
14.5526 | -34.7834 | -23.7186
14.5587 | -34.7839 | -23.727
14.5648 | -34.7844 | -23.7352
14.5709 | -34.7849 | -23.7434
14.577 | -34.7853 | -23.7516
14.5831 | -34.7858 | -23.76
14.5892 | -34.7863 | -23.7683
14.5953 | -34.7868 | -23.7764
14.6014 | -34.7873 | -23.7847
14.6075 | -34.7877 | -23.793
14.6136 | -34.7882 | -23.8013
14.6197 | -34.7887 | -23.8096
14.6258 | -34.7892 | -23.8179
14.6319 | -34.7897 | -23.8261
14.638 | -34.7901 | -23.8344
14.6441 | -34.7906 | -23.8427
14.6502 | -34.7911 | -23.8511
14.6563 | -34.7916 | -23.8593
14.6624 | -34.7921 | -23.8676
14.6685 | -34.7926 | -23.8757
14.6746 | -34.7931 | -23.884
14.6807 | -34.7936 | -23.8921
14.6869 | -34.7941 | -23.9004
14.693 | -34.7946 | -23.9086
14.6991 | -34.7951 | -23.9168
14.7052 | -34.7956 | -23.9248
14.7113 | -34.7961 | -23.9332
14.7174 | -34.7966 | -23.9415
14.7235 | -34.7971 | -23.9497
14.7296 | -34.7975 | -23.9579
14.7357 | -34.7981 | -23.9661
14.7418 | -34.7986 | -23.9743
14.7479 | -34.799 | -23.9825
14.754 | -34.7995 | -23.9908
14.7601 | -34.8 | -23.999
14.7662 | -34.8005 | -24.0071
14.7723 | -34.801 | -24.0151
14.7784 | -34.8015 | -24.0234
14.7845 | -34.802 | -24.0316
14.7906 | -34.8025 | -24.04
14.7967 | -34.803 | -24.0481
14.8028 | -34.8034 | -24.0563
14.8089 | -34.804 | -24.0645
14.815 | -34.8045 | -24.0727
14.8211 | -34.805 | -24.0809
14.8272 | -34.8054 | -24.0893
14.8334 | -34.8059 | -24.0975
14.8395 | -34.8064 | -24.1057
14.8456 | -34.8069 | -24.1137
14.8517 | -34.8074 | -24.1219
14.8578 | -34.8079 | -24.1302
14.8639 | -34.8084 | -24.1385
14.87 | -34.8089 | -24.1468
14.8761 | -34.8094 | -24.1549
14.8822 | -34.8099 | -24.1632
14.8883 | -34.8104 | -24.1714
14.8944 | -34.8109 | -24.1796
14.9005 | -34.8114 | -24.1878
14.9066 | -34.8119 | -24.196
14.9127 | -34.8124 | -24.2041
14.9188 | -34.8129 | -24.2123
14.9249 | -34.8134 | -24.2205
14.931 | -34.8139 | -24.2288
14.9371 | -34.8144 | -24.2368
14.9432 | -34.8148 | -24.2451
14.9493 | -34.8153 | -24.2532
14.9554 | -34.8158 | -24.2613
14.9615 | -34.8163 | -24.2695
14.9676 | -34.8168 | -24.2778
14.9738 | -34.8173 | -24.286
14.9799 | -34.8178 | -24.2942
14.986 | -34.8183 | -24.3023
14.9921 | -34.8188 | -24.3104
14.9982 | -34.8192 | -24.3186
15.0043 | -34.8197 | -24.3267
15.0104 | -34.8202 | -24.335
15.0165 | -34.8207 | -24.3433
15.0226 | -34.8213 | -24.3513
15.0287 | -34.8217 | -24.3595
15.0348 | -34.8222 | -24.3677
15.0409 | -34.8227 | -24.3759
15.047 | -34.8232 | -24.3843
15.0531 | -34.8238 | -24.3922
15.0592 | -34.8243 | -24.4004
15.0653 | -34.8248 | -24.4087
15.0714 | -34.8253 | -24.4168
15.0775 | -34.8258 | -24.4248
15.0836 | -34.8263 | -24.433
15.0897 | -34.8268 | -24.4409
15.0958 | -34.8273 | -24.4491
15.1019 | -34.8278 | -24.4572
15.108 | -34.8283 | -24.4653
15.1141 | -34.8288 | -24.4734
15.1203 | -34.8293 | -24.4816
15.1264 | -34.8298 | -24.4897
15.1325 | -34.8303 | -24.4979
15.1386 | -34.8308 | -24.506
15.1447 | -34.8313 | -24.5141
15.1508 | -34.8318 | -24.5222
15.1569 | -34.8323 | -24.5304
15.163 | -34.8328 | -24.5383
15.1691 | -34.8333 | -24.5463
15.1752 | -34.8338 | -24.5544
15.1813 | -34.8343 | -24.5625
15.1874 | -34.8348 | -24.5706
15.1935 | -34.8353 | -24.5786
15.1996 | -34.8358 | -24.5866
15.2057 | -34.8363 | -24.5947
15.2118 | -34.8368 | -24.6028
15.2179 | -34.8373 | -24.611
15.224 | -34.8378 | -24.619
15.2301 | -34.8383 | -24.6269
15.2362 | -34.8388 | -24.635
15.2423 | -34.8393 | -24.6433
15.2484 | -34.8398 | -24.6514
15.2545 | -34.8403 | -24.6596
15.2607 | -34.8408 | -24.6677
15.2668 | -34.8413 | -24.6755
15.2729 | -34.8418 | -24.6835
15.279 | -34.8423 | -24.6917
15.2851 | -34.8428 | -24.6997
15.2912 | -34.8433 | -24.7078
15.2973 | -34.8438 | -24.716
15.3034 | -34.8443 | -24.7239
15.3095 | -34.8448 | -24.7321
15.3156 | -34.8453 | -24.74
15.3217 | -34.8459 | -24.7481
15.3278 | -34.8464 | -24.7563
15.3339 | -34.8469 | -24.7641
15.34 | -34.8474 | -24.7721
15.3461 | -34.8479 | -24.7802
15.3522 | -34.8484 | -24.7882
15.3583 | -34.8489 | -24.7965
15.3644 | -34.8494 | -24.8045
15.3705 | -34.8499 | -24.8122
15.3766 | -34.8504 | -24.8204
15.3827 | -34.8509 | -24.8284
15.3888 | -34.8514 | -24.8365
15.3949 | -34.8519 | -24.8445
15.401 | -34.8524 | -24.8527
15.4072 | -34.8529 | -24.8605
15.4133 | -34.8534 | -24.8686
15.4194 | -34.8539 | -24.8768
15.4255 | -34.8544 | -24.8848
15.4316 | -34.8549 | -24.8929
15.4377 | -34.8554 | -24.901
15.4438 | -34.8559 | -24.9089
15.4499 | -34.8564 | -24.9169
15.456 | -34.8569 | -24.9248
15.4621 | -34.8575 | -24.9328
15.4682 | -34.8579 | -24.9409
15.4743 | -34.8585 | -24.9487
15.4804 | -34.859 | -24.9567
15.4865 | -34.8594 | -24.9647
15.4926 | -34.8599 | -24.9728
15.4987 | -34.8604 | -24.9808
15.5048 | -34.8609 | -24.9888
15.5109 | -34.8614 | -24.9967
15.517 | -34.862 | -25.0048
15.5231 | -34.8625 | -25.0128
15.5292 | -34.8629 | -25.0207
15.5353 | -34.8634 | -25.0288
15.5414 | -34.8639 | -25.0367
15.5476 | -34.8644 | -25.0445
15.5537 | -34.865 | -25.0525
15.5598 | -34.8655 | -25.0606
15.5659 | -34.866 | -25.0685
15.572 | -34.8665 | -25.0764
15.5781 | -34.867 | -25.0843
15.5842 | -34.8675 | -25.0921
15.5903 | -34.868 | -25.1
15.5964 | -34.8685 | -25.1079
15.6025 | -34.869 | -25.1158
15.6086 | -34.8695 | -25.1238
15.6147 | -34.87 | -25.1318
15.6208 | -34.8705 | -25.1395
15.6269 | -34.871 | -25.1474
15.633 | -34.8715 | -25.1555
15.6391 | -34.872 | -25.1635
15.6452 | -34.8725 | -25.1715
15.6513 | -34.8731 | -25.1791
15.6574 | -34.8736 | -25.1872
15.6635 | -34.8741 | -25.1951
15.6696 | -34.8746 | -25.2031
15.6757 | -34.8751 | -25.2111
15.6818 | -34.8756 | -25.2191
15.688 | -34.8761 | -25.2269
15.6941 | -34.8766 | -25.2348
15.7002 | -34.8771 | -25.2427
15.7063 | -34.8775 | -25.2507
15.7124 | -34.878 | -25.2587
15.7185 | -34.8785 | -25.2667
15.7246 | -34.8791 | -25.2745
15.7307 | -34.8796 | -25.2824
15.7368 | -34.8801 | -25.2903
15.7429 | -34.8806 | -25.2982
15.749 | -34.8811 | -25.3062
15.7551 | -34.8816 | -25.3142
15.7612 | -34.8822 | -25.3219
15.7673 | -34.8827 | -25.3298
15.7734 | -34.8832 | -25.3377
15.7795 | -34.8837 | -25.3456
15.7856 | -34.8842 | -25.3533
15.7917 | -34.8847 | -25.361
15.7978 | -34.8852 | -25.3691
15.8039 | -34.8857 | -25.377
15.81 | -34.8862 | -25.3849
15.8161 | -34.8867 | -25.3927
15.8222 | -34.8872 | -25.4005
15.8283 | -34.8877 | -25.4081
15.8345 | -34.8882 | -25.416
15.8406 | -34.8887 | -25.4239
15.8467 | -34.8893 | -25.4319
15.8528 | -34.8898 | -25.4398
15.8589 | -34.8903 | -25.4477
15.865 | -34.8908 | -25.4552
15.8711 | -34.8913 | -25.4632
15.8772 | -34.8918 | -25.471
15.8833 | -34.8923 | -25.479
15.8894 | -34.8929 | -25.4869
15.8955 | -34.8934 | -25.4947
15.9016 | -34.8939 | -25.5024
15.9077 | -34.8943 | -25.5103
15.9138 | -34.8948 | -25.5181
15.9199 | -34.8953 | -25.526
15.926 | -34.8958 | -25.5339
15.9321 | -34.8964 | -25.5416
15.9382 | -34.8969 | -25.5493
15.9443 | -34.8974 | -25.5572
15.9504 | -34.8979 | -25.5651
15.9565 | -34.8984 | -25.5729
15.9626 | -34.8989 | -25.5808
15.9687 | -34.8994 | -25.5885
15.9749 | -34.8999 | -25.5962
15.981 | -34.9004 | -25.6041
15.9871 | -34.9009 | -25.612
15.9932 | -34.9014 | -25.6197
15.9993 | -34.9019 | -25.6275
16.0054 | -34.9024 | -25.635
16.0115 | -34.9029 | -25.6428
16.0176 | -34.9034 | -25.6508
16.0237 | -34.904 | -25.6587
16.0298 | -34.9044 | -25.6665
16.0359 | -34.9049 | -25.6743
16.042 | -34.9055 | -25.6819
16.0481 | -34.906 | -25.6897
16.0542 | -34.9065 | -25.6976
16.0603 | -34.907 | -25.7054
16.0664 | -34.9075 | -25.7131
16.0725 | -34.908 | -25.7207
16.0786 | -34.9085 | -25.7285
16.0847 | -34.909 | -25.7364
16.0908 | -34.9095 | -25.7442
16.0969 | -34.91 | -25.7522
16.103 | -34.9105 | -25.76
16.1091 | -34.9111 | -25.7676
16.1152 | -34.9116 | -25.7753
16.1214 | -34.912 | -25.7831
16.1275 | -34.9126 | -25.7909
16.1336 | -34.9131 | -25.7986
16.1397 | -34.9136 | -25.8064
16.1458 | -34.9141 | -25.8141
16.1519 | -34.9146 | -25.822
16.158 | -34.9151 | -25.8299
16.1641 | -34.9156 | -25.8377
16.1702 | -34.9161 | -25.8455
16.1763 | -34.9166 | -25.8533
16.1824 | -34.9171 | -25.8609
16.1885 | -34.9176 | -25.8686
16.1946 | -34.9181 | -25.8764
16.2007 | -34.9186 | -25.8842
16.2068 | -34.9192 | -25.8919
16.2129 | -34.9197 | -25.8996
16.219 | -34.9202 | -25.9071
16.2251 | -34.9207 | -25.9148
16.2312 | -34.9212 | -25.9226
16.2373 | -34.9217 | -25.9302
16.2434 | -34.9222 | -25.9378
16.2495 | -34.9227 | -25.9453
16.2556 | -34.9232 | -25.9531
16.2618 | -34.9238 | -25.9608
16.2679 | -34.9243 | -25.9686
16.274 | -34.9248 | -25.9764
16.2801 | -34.9253 | -25.984
16.2862 | -34.9258 | -25.9917
16.2923 | -34.9263 | -25.9995
16.2984 | -34.9268 | -26.0073
16.3045 | -34.9273 | -26.015
16.3106 | -34.9278 | -26.0227
16.3167 | -34.9283 | -26.0304
16.3228 | -34.9288 | -26.038
16.3289 | -34.9293 | -26.0458
16.335 | -34.9298 | -26.0535
16.3411 | -34.9303 | -26.0612
16.3472 | -34.9308 | -26.069
16.3533 | -34.9313 | -26.0767
16.3594 | -34.9318 | -26.0842
16.3655 | -34.9323 | -26.0919
16.3716 | -34.9328 | -26.0995
16.3777 | -34.9333 | -26.1072
16.3838 | -34.9338 | -26.1149
16.3899 | -34.9344 | -26.1225
16.396 | -34.9349 | -26.1302
16.4021 | -34.9354 | -26.1377
16.4083 | -34.9359 | -26.1454
16.4144 | -34.9364 | -26.1531
16.4205 | -34.9369 | -26.1608
16.4266 | -34.9374 | -26.1683
16.4327 | -34.9379 | -26.1759
16.4388 | -34.9384 | -26.1837
16.4449 | -34.9389 | -26.1915
16.451 | -34.9394 | -26.1993
16.4571 | -34.9399 | -26.2071
16.4632 | -34.9404 | -26.2146
16.4693 | -34.9409 | -26.2223
16.4754 | -34.9414 | -26.2302
16.4815 | -34.9419 | -26.2378
16.4876 | -34.9424 | -26.2454
16.4937 | -34.9429 | -26.253
16.4998 | -34.9434 | -26.2605
16.5059 | -34.9439 | -26.2682
16.512 | -34.9444 | -26.276
16.5181 | -34.9449 | -26.2838
16.5242 | -34.9454 | -26.2916
16.5303 | -34.946 | -26.2991
16.5364 | -34.9465 | -26.3068
16.5425 | -34.947 | -26.3145
16.5487 | -34.9474 | -26.3221
16.5548 | -34.9479 | -26.3298
16.5609 | -34.9485 | -26.3375
16.567 | -34.949 | -26.3449
16.5731 | -34.9495 | -26.3527
16.5792 | -34.95 | -26.3604
16.5853 | -34.9505 | -26.3681
16.5914 | -34.951 | -26.3758
16.5975 | -34.9515 | -26.3835
16.6036 | -34.952 | -26.391
16.6097 | -34.9525 | -26.3986
16.6158 | -34.953 | -26.4063
16.6219 | -34.9535 | -26.4139
16.628 | -34.954 | -26.4215
16.6341 | -34.9545 | -26.4292
16.6402 | -34.955 | -26.4366
16.6463 | -34.9555 | -26.4443
16.6524 | -34.956 | -26.4517
16.6585 | -34.9565 | -26.4595
16.6646 | -34.957 | -26.467
16.6707 | -34.9575 | -26.4745
16.6768 | -34.958 | -26.4823
16.6829 | -34.9586 | -26.4898
16.689 | -34.959 | -26.4975
16.6952 | -34.9595 | -26.5052
16.7013 | -34.96 | -26.5128
16.7074 | -34.9605 | -26.5202
16.7135 | -34.961 | -26.5279
16.7196 | -34.9615 | -26.5354
16.7257 | -34.962 | -26.5431
16.7318 | -34.9625 | -26.5507
16.7379 | -34.963 | -26.5582
16.744 | -34.9635 | -26.5657
16.7501 | -34.964 | -26.5733
16.7562 | -34.9645 | -26.581
16.7623 | -34.965 | -26.5886
16.7684 | -34.9655 | -26.5963
16.7745 | -34.966 | -26.6039
16.7806 | -34.9665 | -26.6113
16.7867 | -34.967 | -26.619
16.7928 | -34.9675 | -26.6266
16.7989 | -34.968 | -26.634
16.805 | -34.9685 | -26.6418
16.8111 | -34.969 | -26.6492
16.8172 | -34.9695 | -26.657
16.8233 | -34.97 | -26.6645
16.8294 | -34.9705 | -26.6721
16.8356 | -34.971 | -26.6797
16.8417 | -34.9715 | -26.6873
16.8478 | -34.972 | -26.6948
16.8539 | -34.9725 | -26.7024
16.86 | -34.973 | -26.7101
16.8661 | -34.9735 | -26.7177
16.8722 | -34.974 | -26.7254
16.8783 | -34.9745 | -26.733
16.8844 | -34.975 | -26.7405
16.8905 | -34.9755 | -26.748
16.8966 | -34.9761 | -26.7556
16.9027 | -34.9766 | -26.7632
16.9088 | -34.9771 | -26.7709
16.9149 | -34.9776 | -26.7785
16.921 | -34.9781 | -26.7859
16.9271 | -34.9786 | -26.7936
16.9332 | -34.9791 | -26.8012
16.9393 | -34.9796 | -26.8089
16.9454 | -34.9801 | -26.8163
16.9515 | -34.9806 | -26.8239
16.9576 | -34.9811 | -26.8314
16.9637 | -34.9816 | -26.8387
16.9698 | -34.9821 | -26.8462
16.9759 | -34.9826 | -26.8538
16.9821 | -34.9831 | -26.8615
16.9882 | -34.9836 | -26.869
16.9943 | -34.9841 | -26.8763
17.0004 | -34.9846 | -26.8839
17.0065 | -34.9851 | -26.8916
17.0126 | -34.9857 | -26.8991
17.0187 | -34.9862 | -26.9069
17.0248 | -34.9867 | -26.9144
17.0309 | -34.9872 | -26.9219
17.037 | -34.9877 | -26.9293
17.0431 | -34.9882 | -26.9369
17.0492 | -34.9886 | -26.9446
17.0553 | -34.9892 | -26.9522
17.0614 | -34.9897 | -26.9597
17.0675 | -34.9902 | -26.9672
17.0736 | -34.9906 | -26.9748
17.0797 | -34.9911 | -26.9823
17.0858 | -34.9916 | -26.9899
17.0919 | -34.9922 | -26.9975
17.098 | -34.9927 | -27.005
17.1041 | -34.9932 | -27.0126
17.1102 | -34.9937 | -27.02
17.1163 | -34.9942 | -27.0276
17.1225 | -34.9947 | -27.0351
17.1286 | -34.9952 | -27.0425
17.1347 | -34.9957 | -27.05
17.1408 | -34.9962 | -27.0576
17.1469 | -34.9967 | -27.0651
17.153 | -34.9972 | -27.0727
17.1591 | -34.9977 | -27.0802
17.1652 | -34.9982 | -27.0876
17.1713 | -34.9987 | -27.095
17.1774 | -34.9992 | -27.1024
17.1835 | -34.9997 | -27.11
17.1896 | -35.0002 | -27.1174
17.1957 | -35.0007 | -27.1249
17.2018 | -35.0012 | -27.1322
17.2079 | -35.0017 | -27.1399
17.214 | -35.0022 | -27.1476
17.2201 | -35.0027 | -27.1549
17.2262 | -35.0032 | -27.1625
17.2323 | -35.0037 | -27.1701
17.2384 | -35.0042 | -27.1775
17.2445 | -35.0047 | -27.1848
17.2506 | -35.0052 | -27.1924
17.2567 | -35.0057 | -27.2001
17.2628 | -35.0062 | -27.2076
17.269 | -35.0067 | -27.2151
17.2751 | -35.0072 | -27.2227
17.2812 | -35.0077 | -27.2302
17.2873 | -35.0082 | -27.2377
17.2934 | -35.0087 | -27.2453
17.2995 | -35.0092 | -27.2528
17.3056 | -35.0097 | -27.2602
17.3117 | -35.0102 | -27.2678
17.3178 | -35.0107 | -27.2753
17.3239 | -35.0112 | -27.2829
17.33 | -35.0117 | -27.2904
17.3361 | -35.0122 | -27.2978
17.3422 | -35.0127 | -27.3051
17.3483 | -35.0132 | -27.3126
17.3544 | -35.0137 | -27.3202
17.3605 | -35.0141 | -27.3275
17.3666 | -35.0147 | -27.3349
17.3727 | -35.0152 | -27.3425
17.3788 | -35.0156 | -27.3498
17.3849 | -35.0161 | -27.3572
17.391 | -35.0166 | -27.3648
17.3971 | -35.0171 | -27.3723
17.4032 | -35.0176 | -27.3797
17.4094 | -35.0181 | -27.3871
17.4155 | -35.0186 | -27.3945
17.4216 | -35.0191 | -27.4018
17.4277 | -35.0196 | -27.4093
17.4338 | -35.0201 | -27.4167
17.4399 | -35.0206 | -27.4243
17.446 | -35.0211 | -27.4316
17.4521 | -35.0216 | -27.4391
17.4582 | -35.0221 | -27.4464
17.4643 | -35.0226 | -27.4539
17.4704 | -35.0231 | -27.4613
17.4765 | -35.0236 | -27.4689
17.4826 | -35.0241 | -27.4762
17.4887 | -35.0246 | -27.4838
17.4948 | -35.0251 | -27.4912
17.5009 | -35.0257 | -27.4987
17.507 | -35.0262 | -27.5062
17.5131 | -35.0267 | -27.5136
17.5192 | -35.0272 | -27.521
17.5253 | -35.0277 | -27.5285
17.5314 | -35.0281 | -27.5358
17.5375 | -35.0286 | -27.5434
17.5436 | -35.0291 | -27.5509
17.5497 | -35.0296 | -27.5585
17.5559 | -35.0301 | -27.5657
17.562 | -35.0307 | -27.5731
17.5681 | -35.0312 | -27.5806
17.5742 | -35.0316 | -27.5881
17.5803 | -35.0322 | -27.5955
17.5864 | -35.0327 | -27.6029
17.5925 | -35.0332 | -27.6105
17.5986 | -35.0337 | -27.6179
17.6047 | -35.0342 | -27.6254
17.6108 | -35.0347 | -27.6328
17.6169 | -35.0352 | -27.6402
17.623 | -35.0357 | -27.6475
17.6291 | -35.0362 | -27.6548
17.6352 | -35.0366 | -27.6622
17.6413 | -35.0371 | -27.6696
17.6474 | -35.0377 | -27.6771
17.6535 | -35.0382 | -27.6845
17.6596 | -35.0387 | -27.6918
17.6657 | -35.0392 | -27.6993
17.6718 | -35.0396 | -27.7066
17.6779 | -35.0401 | -27.714
17.684 | -35.0406 | -27.7214
17.6901 | -35.0411 | -27.7288
17.6963 | -35.0416 | -27.7361
17.7024 | -35.0421 | -27.7434
17.7085 | -35.0426 | -27.7507
17.7146 | -35.0431 | -27.7581
17.7207 | -35.0436 | -27.7654
17.7268 | -35.0442 | -27.7727
17.7329 | -35.0446 | -27.7803
17.739 | -35.0451 | -27.7878
17.7451 | -35.0456 | -27.7951
17.7512 | -35.0461 | -27.8026
17.7573 | -35.0467 | -27.8101
17.7634 | -35.0472 | -27.8173
17.7695 | -35.0477 | -27.8247
17.7756 | -35.0482 | -27.832
17.7817 | -35.0486 | -27.8394
17.7878 | -35.0491 | -27.8468
17.7939 | -35.0496 | -27.8542
17.8 | -35.0501 | -27.8614
17.8061 | -35.0507 | -27.8688
17.8122 | -35.0512 | -27.8761
17.8183 | -35.0517 | -27.8836
17.8244 | -35.0521 | -27.891
17.8305 | -35.0526 | -27.8984
17.8366 | -35.0531 | -27.9056
17.8428 | -35.0536 | -27.9131
17.8489 | -35.0541 | -27.9205
17.855 | -35.0546 | -27.9278
17.8611 | -35.0551 | -27.9351
17.8672 | -35.0556 | -27.9422
17.8733 | -35.0561 | -27.9495
17.8794 | -35.0566 | -27.9568
17.8855 | -35.0571 | -27.9643
17.8916 | -35.0576 | -27.9717
17.8977 | -35.0581 | -27.9791
17.9038 | -35.0586 | -27.9863
17.9099 | -35.0591 | -27.9936
17.916 | -35.0596 | -28.0011
17.9221 | -35.0601 | -28.0084
17.9282 | -35.0606 | -28.0158
17.9343 | -35.0611 | -28.0233
17.9404 | -35.0616 | -28.0304
17.9465 | -35.0621 | -28.0379
17.9526 | -35.0626 | -28.0453
17.9587 | -35.063 | -28.0527
17.9648 | -35.0635 | -28.0599
17.9709 | -35.064 | -28.0673
17.977 | -35.0645 | -28.0747
17.9832 | -35.065 | -28.0821
17.9893 | -35.0655 | -28.0894
17.9954 | -35.066 | -28.0968
18.0015 | -35.0665 | -28.1041
18.0076 | -35.067 | -28.1113
18.0137 | -35.0675 | -28.1186
18.0198 | -35.068 | -28.1259
18.0259 | -35.0685 | -28.1334
18.032 | -35.069 | -28.1408
18.0381 | -35.0695 | -28.1483
18.0442 | -35.07 | -28.1556
18.0503 | -35.0705 | -28.1629
18.0564 | -35.071 | -28.1702
18.0625 | -35.0715 | -28.1776
18.0686 | -35.0719 | -28.185
18.0747 | -35.0724 | -28.1924
18.0808 | -35.0729 | -28.1995
18.0869 | -35.0734 | -28.2067
18.093 | -35.0739 | -28.2141
18.0991 | -35.0744 | -28.2215
18.1052 | -35.0749 | -28.2288
18.1113 | -35.0754 | -28.2361
18.1174 | -35.0759 | -28.2433
18.1236 | -35.0764 | -28.2507
18.1297 | -35.0769 | -28.258
18.1358 | -35.0774 | -28.2654
18.1419 | -35.0779 | -28.2727
18.148 | -35.0784 | -28.28
18.1541 | -35.0789 | -28.2871
18.1602 | -35.0794 | -28.2945
18.1663 | -35.0799 | -28.3018
18.1724 | -35.0804 | -28.3092
18.1785 | -35.0809 | -28.3165
18.1846 | -35.0814 | -28.3237
18.1907 | -35.0819 | -28.331
18.1968 | -35.0824 | -28.3384
18.2029 | -35.0829 | -28.3456
18.209 | -35.0834 | -28.3531
18.2151 | -35.0839 | -28.3605
18.2212 | -35.0844 | -28.3678
18.2273 | -35.0849 | -28.3751
18.2334 | -35.0854 | -28.3824
18.2395 | -35.0859 | -28.3897
18.2456 | -35.0864 | -28.397
18.2517 | -35.0869 | -28.4043
18.2578 | -35.0874 | -28.4116
18.2639 | -35.0879 | -28.4189
18.2701 | -35.0884 | -28.4263
18.2762 | -35.0888 | -28.4336
18.2823 | -35.0893 | -28.4408
18.2884 | -35.0898 | -28.448
18.2945 | -35.0903 | -28.455
18.3006 | -35.0908 | -28.4624
18.3067 | -35.0913 | -28.4696
18.3128 | -35.0918 | -28.477
18.3189 | -35.0923 | -28.4843
18.325 | -35.0928 | -28.4914
18.3311 | -35.0933 | -28.4987
18.3372 | -35.0938 | -28.506
18.3433 | -35.0943 | -28.5133
18.3494 | -35.0948 | -28.5205
18.3555 | -35.0953 | -28.5278
18.3616 | -35.0958 | -28.5349
18.3677 | -35.0963 | -28.5422
18.3738 | -35.0968 | -28.5494
18.3799 | -35.0973 | -28.5568
18.386 | -35.0978 | -28.564
18.3921 | -35.0983 | -28.5712
18.3982 | -35.0988 | -28.5782
18.4043 | -35.0992 | -28.5855
18.4105 | -35.0997 | -28.5928
18.4166 | -35.1002 | -28.6001
18.4227 | -35.1007 | -28.6074
18.4288 | -35.1012 | -28.6147
18.4349 | -35.1017 | -28.6219
18.441 | -35.1021 | -28.629
18.4471 | -35.1026 | -28.6362
18.4532 | -35.1031 | -28.6435
18.4593 | -35.1036 | -28.6508
18.4654 | -35.1041 | -28.6579
18.4715 | -35.1046 | -28.6653
18.4776 | -35.105 | -28.6727
18.4837 | -35.1055 | -28.68
18.4898 | -35.106 | -28.6873
18.4959 | -35.1065 | -28.6947
18.502 | -35.107 | -28.702
18.5081 | -35.1075 | -28.7092
18.5142 | -35.108 | -28.7165
18.5203 | -35.1085 | -28.7235
18.5264 | -35.109 | -28.7307
18.5325 | -35.1095 | -28.738
18.5386 | -35.11 | -28.7451
18.5447 | -35.1105 | -28.7523
18.5508 | -35.111 | -28.7594
18.557 | -35.1115 | -28.7666
18.5631 | -35.1119 | -28.774
18.5692 | -35.1124 | -28.7812
18.5753 | -35.1129 | -28.7882
18.5814 | -35.1134 | -28.7954
18.5875 | -35.1139 | -28.8025
18.5936 | -35.1144 | -28.8097
18.5997 | -35.1149 | -28.817
18.6058 | -35.1154 | -28.8242
18.6119 | -35.1159 | -28.8314
18.618 | -35.1164 | -28.8387
18.6241 | -35.1168 | -28.846
18.6302 | -35.1173 | -28.8534
18.6363 | -35.1178 | -28.8606
18.6424 | -35.1183 | -28.8676
18.6485 | -35.1188 | -28.8749
18.6546 | -35.1193 | -28.8821
18.6607 | -35.1198 | -28.8893
18.6668 | -35.1203 | -28.8965
18.6729 | -35.1208 | -28.9039
18.679 | -35.1213 | -28.9109
18.6851 | -35.1218 | -28.9179
18.6912 | -35.1223 | -28.9252
18.6974 | -35.1228 | -28.9323
18.7035 | -35.1233 | -28.9395
18.7096 | -35.1238 | -28.9468
18.7157 | -35.1242 | -28.954
18.7218 | -35.1247 | -28.9612
18.7279 | -35.1253 | -28.9684
18.734 | -35.1258 | -28.9756
18.7401 | -35.1263 | -28.9828
18.7462 | -35.1268 | -28.9896
18.7523 | -35.1273 | -28.9968
18.7584 | -35.1278 | -29.0042
18.7645 | -35.1283 | -29.0113
18.7706 | -35.1287 | -29.0184
18.7767 | -35.1292 | -29.0256
18.7828 | -35.1297 | -29.0325
18.7889 | -35.1302 | -29.0398
18.795 | -35.1307 | -29.047
18.8011 | -35.1312 | -29.0541
18.8072 | -35.1317 | -29.0614
18.8133 | -35.1321 | -29.0686
18.8194 | -35.1326 | -29.0756
18.8255 | -35.1331 | -29.0828
18.8316 | -35.1336 | -29.0899
18.8377 | -35.1341 | -29.0971
18.8439 | -35.1346 | -29.1043
18.85 | -35.1351 | -29.1115
18.8561 | -35.1356 | -29.1185
18.8622 | -35.136 | -29.1259
18.8683 | -35.1365 | -29.133
18.8744 | -35.137 | -29.1401
18.8805 | -35.1375 | -29.1474
18.8866 | -35.138 | -29.1547
18.8927 | -35.1385 | -29.1618
18.8988 | -35.139 | -29.1688
18.9049 | -35.1395 | -29.176
18.911 | -35.14 | -29.183
18.9171 | -35.1405 | -29.1901
18.9232 | -35.141 | -29.197
18.9293 | -35.1415 | -29.2042
18.9354 | -35.1419 | -29.2112
18.9415 | -35.1424 | -29.2184
18.9476 | -35.1429 | -29.2255
18.9537 | -35.1434 | -29.2327
18.9598 | -35.1439 | -29.2396
18.9659 | -35.1444 | -29.2468
18.972 | -35.1449 | -29.2538
18.9781 | -35.1453 | -29.2609
18.9843 | -35.1458 | -29.2682
18.9904 | -35.1463 | -29.2754
18.9965 | -35.1468 | -29.2824
19.0026 | -35.1473 | -29.2896
19.0087 | -35.1477 | -29.2969
19.0148 | -35.1483 | -29.304
19.0209 | -35.1487 | -29.3112
19.027 | -35.1493 | -29.3183
19.0331 | -35.1497 | -29.3254
19.0392 | -35.1502 | -29.3324
19.0453 | -35.1507 | -29.3396
19.0514 | -35.1512 | -29.3466
19.0575 | -35.1517 | -29.3539
19.0636 | -35.1522 | -29.3608
19.0697 | -35.1527 | -29.3679
19.0758 | -35.1532 | -29.3749
19.0819 | -35.1536 | -29.3819
19.088 | -35.1541 | -29.389
19.0941 | -35.1546 | -29.3962
19.1002 | -35.1551 | -29.4033
19.1063 | -35.1556 | -29.4104
19.1124 | -35.1561 | -29.4175
19.1185 | -35.1565 | -29.4246
19.1246 | -35.157 | -29.4318
19.1308 | -35.1575 | -29.439
19.1369 | -35.158 | -29.4461
19.143 | -35.1585 | -29.4532
19.1491 | -35.159 | -29.4604
19.1552 | -35.1595 | -29.4676
19.1613 | -35.16 | -29.4748
19.1674 | -35.1605 | -29.4819
19.1735 | -35.161 | -29.4888
19.1796 | -35.1615 | -29.4958
19.1857 | -35.1619 | -29.503
19.1918 | -35.1624 | -29.5101
19.1979 | -35.163 | -29.5172
19.204 | -35.1634 | -29.5242
19.2101 | -35.1639 | -29.5312
19.2162 | -35.1644 | -29.5382
19.2223 | -35.1649 | -29.5454
19.2284 | -35.1654 | -29.5526
19.2345 | -35.1658 | -29.5598
19.2406 | -35.1663 | -29.5667
19.2467 | -35.1668 | -29.5738
19.2528 | -35.1673 | -29.5809
19.2589 | -35.1678 | -29.5878
19.265 | -35.1683 | -29.5949
19.2712 | -35.1687 | -29.602
19.2773 | -35.1692 | -29.6089
19.2834 | -35.1697 | -29.616
19.2895 | -35.1702 | -29.6231
19.2956 | -35.1707 | -29.6301
19.3017 | -35.1712 | -29.6371
19.3078 | -35.1717 | -29.6443
19.3139 | -35.1721 | -29.6512
19.32 | -35.1726 | -29.6581
19.3261 | -35.1731 | -29.6651
19.3322 | -35.1736 | -29.6721
19.3383 | -35.1741 | -29.6791
19.3444 | -35.1746 | -29.686
19.3505 | -35.1751 | -29.6929
19.3566 | -35.1756 | -29.6999
19.3627 | -35.176 | -29.7069
19.3688 | -35.1765 | -29.7139
19.3749 | -35.177 | -29.721
19.381 | -35.1775 | -29.7277
19.3871 | -35.178 | -29.7346
19.3932 | -35.1785 | -29.7417
19.3993 | -35.179 | -29.7487
19.4054 | -35.1794 | -29.7559
19.4115 | -35.1799 | -29.763
19.4177 | -35.1804 | -29.7698
19.4238 | -35.1809 | -29.7768
19.4299 | -35.1814 | -29.7837
19.436 | -35.1818 | -29.7908
19.4421 | -35.1823 | -29.7976
19.4482 | -35.1828 | -29.8046
19.4543 | -35.1833 | -29.8114
19.4604 | -35.1837 | -29.8184
19.4665 | -35.1842 | -29.8253
19.4726 | -35.1847 | -29.8322
19.4787 | -35.1852 | -29.8392
19.4848 | -35.1856 | -29.8461
19.4909 | -35.1861 | -29.853
19.497 | -35.1866 | -29.8599
19.5031 | -35.1871 | -29.8668
19.5092 | -35.1875 | -29.8739
19.5153 | -35.188 | -29.8808
19.5214 | -35.1885 | -29.8876
19.5275 | -35.189 | -29.8944
19.5336 | -35.1895 | -29.9014
19.5397 | -35.1899 | -29.9084
19.5458 | -35.1904 | -29.9155
19.5519 | -35.1909 | -29.9223
19.5581 | -35.1913 | -29.929
19.5642 | -35.1918 | -29.9359
19.5703 | -35.1922 | -29.9429
19.5764 | -35.1927 | -29.9499
19.5825 | -35.1932 | -29.9569
19.5886 | -35.1937 | -29.964
19.5947 | -35.1941 | -29.9708
19.6008 | -35.1946 | -29.9777
19.6069 | -35.1951 | -29.9846
19.613 | -35.1956 | -29.9917
19.6191 | -35.1961 | -29.9987
19.6252 | -35.1965 | -30.0057
19.6313 | -35.197 | -30.0126
19.6374 | -35.1975 | -30.0196
19.6435 | -35.198 | -30.0265
19.6496 | -35.1985 | -30.0332
19.6557 | -35.199 | -30.0402
19.6618 | -35.1994 | -30.0468
19.6679 | -35.1999 | -30.0537
19.674 | -35.2004 | -30.0607
19.6801 | -35.2008 | -30.0676
19.6862 | -35.2013 | -30.0745
19.6923 | -35.2018 | -30.0813
19.6984 | -35.2023 | -30.0881
19.7046 | -35.2028 | -30.095
19.7107 | -35.2033 | -30.1017
19.7168 | -35.2038 | -30.1086
19.7229 | -35.2043 | -30.1156
19.729 | -35.2048 | -30.1225
19.7351 | -35.2052 | -30.129
19.7412 | -35.2057 | -30.1361
19.7473 | -35.2062 | -30.143
19.7534 | -35.2067 | -30.1498
19.7595 | -35.2071 | -30.1567
19.7656 | -35.2076 | -30.1638
19.7717 | -35.2081 | -30.1705
19.7778 | -35.2085 | -30.1773
19.7839 | -35.209 | -30.184
19.79 | -35.2095 | -30.1909
19.7961 | -35.21 | -30.1979
19.8022 | -35.2104 | -30.2045
19.8083 | -35.2109 | -30.2113
19.8144 | -35.2114 | -30.2181
19.8205 | -35.2119 | -30.225
19.8266 | -35.2124 | -30.2318
19.8327 | -35.2129 | -30.2387
19.8388 | -35.2133 | -30.2453
19.845 | -35.2138 | -30.252
19.8511 | -35.2143 | -30.2589
19.8572 | -35.2148 | -30.2656
19.8633 | -35.2152 | -30.2726
19.8694 | -35.2157 | -30.2797
19.8755 | -35.2162 | -30.2863
19.8816 | -35.2167 | -30.2931
19.8877 | -35.2172 | -30.3
19.8938 | -35.2177 | -30.3068
19.8999 | -35.2181 | -30.3136
19.906 | -35.2186 | -30.3206
19.9121 | -35.219 | -30.3271
19.9182 | -35.2195 | -30.334
19.9243 | -35.22 | -30.3408
19.9304 | -35.2205 | -30.3475
19.9365 | -35.2209 | -30.3543
19.9426 | -35.2213 | -30.3609
19.9487 | -35.2218 | -30.3677
19.9548 | -35.2223 | -30.3745
19.9609 | -35.2228 | -30.3813
19.967 | -35.2233 | -30.3882
19.9731 | -35.2237 | -30.3949
19.9792 | -35.2242 | -30.4017
19.9853 | -35.2246 | -30.4084
19.9915 | -35.2251 | -30.4151
19.9976 | -35.2256 | -30.422
20.0037 | -35.2261 | -30.4287
20.0098 | -35.2265 | -30.4355
20.0159 | -35.227 | -30.4423
20.022 | -35.2275 | -30.4492
20.0281 | -35.2279 | -30.456
20.0342 | -35.2284 | -30.4629
20.0403 | -35.2289 | -30.4697
20.0464 | -35.2294 | -30.4764
20.0525 | -35.2298 | -30.4829
20.0586 | -35.2303 | -30.4895
20.0647 | -35.2308 | -30.4963
20.0708 | -35.2313 | -30.5032
20.0769 | -35.2318 | -30.51
20.083 | -35.2323 | -30.5166
20.0891 | -35.2327 | -30.5234
20.0952 | -35.2332 | -30.53
20.1013 | -35.2336 | -30.5369
20.1074 | -35.2341 | -30.5435
20.1135 | -35.2346 | -30.5505
20.1196 | -35.235 | -30.5569
20.1257 | -35.2355 | -30.5636
20.1319 | -35.236 | -30.5704
20.138 | -35.2364 | -30.5773
20.1441 | -35.2369 | -30.5842
20.1502 | -35.2373 | -30.5911
20.1563 | -35.2378 | -30.5976
20.1624 | -35.2382 | -30.6043
20.1685 | -35.2387 | -30.6112
20.1746 | -35.2391 | -30.6179
20.1807 | -35.2396 | -30.6248
20.1868 | -35.2401 | -30.6315
20.1929 | -35.2406 | -30.6378
20.199 | -35.2411 | -30.6447
20.2051 | -35.2415 | -30.6514
20.2112 | -35.242 | -30.6581
20.2173 | -35.2425 | -30.6649
20.2234 | -35.243 | -30.6715
20.2295 | -35.2434 | -30.678
20.2356 | -35.2439 | -30.6849
20.2417 | -35.2443 | -30.6917
20.2478 | -35.2448 | -30.6984
20.2539 | -35.2453 | -30.7053
20.26 | -35.2458 | -30.7118
20.2661 | -35.2462 | -30.7187
20.2723 | -35.2467 | -30.7255
20.2784 | -35.2472 | -30.7323
20.2845 | -35.2477 | -30.7387
20.2906 | -35.2482 | -30.7454
20.2967 | -35.2486 | -30.7518
20.3028 | -35.2491 | -30.7584
20.3089 | -35.2496 | -30.7651
20.315 | -35.2501 | -30.7717
20.3211 | -35.2506 | -30.7786
20.3272 | -35.2511 | -30.7853
20.3333 | -35.2515 | -30.7917
20.3394 | -35.252 | -30.7985
20.3455 | -35.2525 | -30.8052
20.3516 | -35.2529 | -30.8117
20.3577 | -35.2534 | -30.8183
20.3638 | -35.2539 | -30.8248
20.3699 | -35.2544 | -30.8314
20.376 | -35.2548 | -30.8381
20.3821 | -35.2553 | -30.8448
20.3882 | -35.2558 | -30.8516
20.3943 | -35.2562 | -30.8584
20.4004 | -35.2567 | -30.8648
20.4065 | -35.2571 | -30.8713
20.4126 | -35.2576 | -30.878
20.4188 | -35.2581 | -30.8848
20.4249 | -35.2586 | -30.8915
20.431 | -35.259 | -30.8979
20.4371 | -35.2595 | -30.9042
20.4432 | -35.26 | -30.9107
20.4493 | -35.2604 | -30.9175
20.4554 | -35.2609 | -30.9241
20.4615 | -35.2614 | -30.9308
20.4676 | -35.2619 | -30.9376
20.4737 | -35.2623 | -30.9439
20.4798 | -35.2628 | -30.9504
20.4859 | -35.2633 | -30.9571
20.492 | -35.2638 | -30.9635
20.4981 | -35.2642 | -30.9699
20.5042 | -35.2647 | -30.9767
20.5103 | -35.2652 | -30.9829
20.5164 | -35.2656 | -30.9895
20.5225 | -35.266 | -30.996
20.5286 | -35.2665 | -31.0027
20.5347 | -35.267 | -31.0091
20.5408 | -35.2674 | -31.0155
20.5469 | -35.2679 | -31.0221
20.553 | -35.2684 | -31.0288
20.5592 | -35.2688 | -31.0354
20.5653 | -35.2693 | -31.0421
20.5714 | -35.2698 | -31.0488
20.5775 | -35.2702 | -31.0551
20.5836 | -35.2707 | -31.0616
20.5897 | -35.2712 | -31.0683
20.5958 | -35.2716 | -31.0749
20.6019 | -35.2721 | -31.0814
20.608 | -35.2726 | -31.0881
20.6141 | -35.2731 | -31.0946
20.6202 | -35.2735 | -31.1011
20.6263 | -35.274 | -31.1078
20.6324 | -35.2745 | -31.1142
20.6385 | -35.275 | -31.1209
20.6446 | -35.2754 | -31.1273
20.6507 | -35.2758 | -31.1336
20.6568 | -35.2763 | -31.1402
20.6629 | -35.2768 | -31.1471
20.669 | -35.2772 | -31.1537
20.6751 | -35.2777 | -31.1603
20.6812 | -35.2782 | -31.1665
20.6873 | -35.2786 | -31.1732
20.6934 | -35.2791 | -31.1797
20.6995 | -35.2795 | -31.1861
20.7057 | -35.28 | -31.1929
20.7118 | -35.2805 | -31.1995
20.7179 | -35.2809 | -31.2058
20.724 | -35.2814 | -31.2124
20.7301 | -35.2818 | -31.2189
20.7362 | -35.2823 | -31.2254
20.7423 | -35.2828 | -31.232
20.7484 | -35.2833 | -31.2387
20.7545 | -35.2837 | -31.245
20.7606 | -35.2842 | -31.2517
20.7667 | -35.2847 | -31.2584
20.7728 | -35.2851 | -31.2648
20.7789 | -35.2856 | -31.2713
20.785 | -35.2861 | -31.2779
20.7911 | -35.2865 | -31.2844
20.7972 | -35.287 | -31.291
20.8033 | -35.2874 | -31.2975
20.8094 | -35.2879 | -31.3039
20.8155 | -35.2883 | -31.3104
20.8216 | -35.2887 | -31.317
20.8277 | -35.2892 | -31.3231
20.8338 | -35.2897 | -31.3295
20.8399 | -35.2901 | -31.3361
20.8461 | -35.2906 | -31.3425
20.8522 | -35.2911 | -31.3492
20.8583 | -35.2915 | -31.356
20.8644 | -35.292 | -31.3626
20.8705 | -35.2925 | -31.3691
20.8766 | -35.2929 | -31.3755
20.8827 | -35.2934 | -31.3822
20.8888 | -35.2939 | -31.3888
20.8949 | -35.2943 | -31.3954
20.901 | -35.2948 | -31.4021
20.9071 | -35.2952 | -31.4086
20.9132 | -35.2957 | -31.415
20.9193 | -35.2962 | -31.4215
20.9254 | -35.2966 | -31.428
20.9315 | -35.2972 | -31.4341
20.9376 | -35.2976 | -31.4407
20.9437 | -35.2981 | -31.4475
20.9498 | -35.2986 | -31.4541
20.9559 | -35.299 | -31.4607
20.962 | -35.2995 | -31.4674
20.9681 | -35.2999 | -31.4736
20.9742 | -35.3004 | -31.4801
20.9803 | -35.3009 | -31.4869
20.9864 | -35.3013 | -31.4935
20.9926 | -35.3018 | -31.5003
20.9987 | -35.3023 | -31.5065
21.0048 | -35.3028 | -31.5129
21.0109 | -35.3032 | -31.5196
21.017 | -35.3037 | -31.526
21.0231 | -35.3042 | -31.5323
21.0292 | -35.3046 | -31.5386
21.0353 | -35.3051 | -31.545
21.0414 | -35.3055 | -31.5515
21.0475 | -35.3059 | -31.558
21.0536 | -35.3064 | -31.5646
21.0597 | -35.3069 | -31.571
21.0658 | -35.3074 | -31.577
21.0719 | -35.3078 | -31.5837
21.078 | -35.3083 | -31.5902
21.0841 | -35.3087 | -31.5969
21.0902 | -35.3092 | -31.6033
21.0963 | -35.3097 | -31.6096
21.1024 | -35.3103 | -31.6164
21.1085 | -35.3108 | -31.6225
21.1146 | -35.3113 | -31.6291
21.1207 | -35.3117 | -31.6355
21.1268 | -35.3122 | -31.6418
21.133 | -35.3127 | -31.6484
21.1391 | -35.3131 | -31.6543
21.1452 | -35.3136 | -31.6607
21.1513 | -35.3141 | -31.6671
21.1574 | -35.3146 | -31.6736
21.1635 | -35.3151 | -31.6796
21.1696 | -35.3155 | -31.6862
21.1757 | -35.316 | -31.6929
21.1818 | -35.3165 | -31.6993
21.1879 | -35.3169 | -31.7056
21.194 | -35.3174 | -31.7119
21.2001 | -35.3179 | -31.7184
21.2062 | -35.3183 | -31.7252
21.2123 | -35.3188 | -31.7317
21.2184 | -35.3193 | -31.7384
21.2245 | -35.3197 | -31.7449
21.2306 | -35.3202 | -31.7513
21.2367 | -35.3206 | -31.7581
21.2428 | -35.321 | -31.7646
21.2489 | -35.3216 | -31.7715
21.255 | -35.3221 | -31.7778
21.2611 | -35.3225 | -31.7843
21.2672 | -35.3229 | -31.7906
21.2733 | -35.3234 | -31.7971
21.2795 | -35.3237 | -31.8035
21.2856 | -35.3242 | -31.8097
21.2917 | -35.3247 | -31.8166
21.2978 | -35.3251 | -31.823
21.3039 | -35.3256 | -31.8294
21.31 | -35.326 | -31.8361
21.3161 | -35.3264 | -31.8427
21.3222 | -35.3268 | -31.8491
21.3283 | -35.3273 | -31.8553
21.3344 | -35.3277 | -31.8619
21.3405 | -35.3282 | -31.8684
21.3466 | -35.3286 | -31.8748
21.3527 | -35.3291 | -31.8812
21.3588 | -35.3295 | -31.8873
21.3649 | -35.3299 | -31.8937
21.371 | -35.3304 | -31.9001
21.3771 | -35.3309 | -31.9064
21.3832 | -35.3313 | -31.9128
21.3893 | -35.3318 | -31.9194
21.3954 | -35.3323 | -31.9258
21.4015 | -35.3327 | -31.9318
21.4076 | -35.3331 | -31.9378
21.4137 | -35.3336 | -31.9444
21.4199 | -35.3341 | -31.9509
21.426 | -35.3346 | -31.9574
21.4321 | -35.335 | -31.9634
21.4382 | -35.3354 | -31.9697
21.4443 | -35.3358 | -31.9761
21.4504 | -35.3363 | -31.9827
21.4565 | -35.3367 | -31.9895
21.4626 | -35.3371 | -31.9956
21.4687 | -35.3376 | -32.0019
21.4748 | -35.3381 | -32.0076
21.4809 | -35.3385 | -32.0141
21.487 | -35.3389 | -32.0204
21.4931 | -35.3394 | -32.0265
21.4992 | -35.3399 | -32.0328
21.5053 | -35.3404 | -32.0388
21.5114 | -35.3408 | -32.0455
21.5175 | -35.3413 | -32.0515
21.5236 | -35.3417 | -32.0577
21.5297 | -35.3423 | -32.0636
21.5358 | -35.3427 | -32.0698
21.5419 | -35.3432 | -32.0762
21.548 | -35.3437 | -32.0826
21.5541 | -35.3442 | -32.0886
21.5602 | -35.3446 | -32.0951
21.5664 | -35.345 | -32.101
21.5725 | -35.3454 | -32.1071
21.5786 | -35.3458 | -32.1136
21.5847 | -35.3463 | -32.1201
21.5908 | -35.3467 | -32.1264
21.5969 | -35.3472 | -32.1325
21.603 | -35.3475 | -32.1391
21.6091 | -35.348 | -32.1457
21.6152 | -35.3484 | -32.1521
21.6213 | -35.3489 | -32.1583
21.6274 | -35.3493 | -32.1648
21.6335 | -35.3498 | -32.1704
21.6396 | -35.3502 | -32.177
21.6457 | -35.3507 | -32.1834
21.6518 | -35.3511 | -32.1898
21.6579 | -35.3516 | -32.1962
21.664 | -35.3521 | -32.2028
21.6701 | -35.3525 | -32.2092
21.6762 | -35.3529 | -32.2157
21.6823 | -35.3534 | -32.2223
21.6884 | -35.3539 | -32.2287
21.6945 | -35.3543 | -32.2349
21.7006 | -35.3548 | -32.2412
21.7068 | -35.3552 | -32.2469
21.7129 | -35.3557 | -32.2528
21.719 | -35.3561 | -32.2593
21.7251 | -35.3566 | -32.2655
21.7312 | -35.3571 | -32.2717
21.7373 | -35.3576 | -32.2774
21.7434 | -35.3581 | -32.2839
21.7495 | -35.3585 | -32.29
21.7556 | -35.359 | -32.2962
21.7617 | -35.3594 | -32.3024
21.7678 | -35.3599 | -32.3087
21.7739 | -35.3603 | -32.3157
21.78 | -35.3608 | -32.3216
21.7861 | -35.3612 | -32.3276
21.7922 | -35.3616 | -32.3337
21.7983 | -35.362 | -32.3398
21.8044 | -35.3625 | -32.3461
21.8105 | -35.363 | -32.3523
21.8166 | -35.3635 | -32.3583
21.8227 | -35.3639 | -32.3642
21.8288 | -35.3644 | -32.3706
21.8349 | -35.3647 | -32.3772
21.841 | -35.3651 | -32.3835
21.8471 | -35.3657 | -32.3898
21.8533 | -35.3661 | -32.3958
21.8594 | -35.3665 | -32.402
21.8655 | -35.3669 | -32.4086
21.8716 | -35.3673 | -32.4148
21.8777 | -35.3678 | -32.4208
21.8838 | -35.3683 | -32.4266
21.8899 | -35.3688 | -32.4329
21.896 | -35.3692 | -32.4386
21.9021 | -35.3697 | -32.4452
21.9082 | -35.3701 | -32.4514
21.9143 | -35.3705 | -32.4576
21.9204 | -35.3709 | -32.4639
21.9265 | -35.3714 | -32.4702
21.9326 | -35.3718 | -32.4762
21.9387 | -35.3722 | -32.4823
21.9448 | -35.3726 | -32.4889
21.9509 | -35.3731 | -32.4948
21.957 | -35.3734 | -32.5011
21.9631 | -35.374 | -32.5068
21.9692 | -35.3744 | -32.5129
21.9753 | -35.3749 | -32.5194
21.9814 | -35.3753 | -32.5253
21.9875 | -35.3757 | -32.5313
21.9937 | -35.3763 | -32.5376
21.9998 | -35.3767 | -32.5438
22.0059 | -35.3772 | -32.5499
22.012 | -35.3777 | -32.556
22.0181 | -35.3781 | -32.5617
22.0242 | -35.3786 | -32.568
22.0303 | -35.3791 | -32.5744
22.0364 | -35.3795 | -32.5808
22.0425 | -35.3799 | -32.5873
22.0486 | -35.3805 | -32.5934
22.0547 | -35.3809 | -32.5993
22.0608 | -35.3814 | -32.6054
22.0669 | -35.3819 | -32.6116
22.073 | -35.3823 | -32.6177
22.0791 | -35.3827 | -32.6238
22.0852 | -35.3832 | -32.6303
22.0913 | -35.3837 | -32.6361
22.0974 | -35.3842 | -32.6421
22.1035 | -35.3846 | -32.6482
22.1096 | -35.385 | -32.6544
22.1157 | -35.3855 | -32.6611
22.1218 | -35.3859 | -32.6673
22.1279 | -35.3864 | -32.6732
22.134 | -35.3868 | -32.6795
22.1402 | -35.3873 | -32.6858
22.1463 | -35.3877 | -32.692
22.1524 | -35.3881 | -32.6983
22.1585 | -35.3886 | -32.7046
22.1646 | -35.3891 | -32.7108
22.1707 | -35.3895 | -32.7169
22.1768 | -35.3899 | -32.7232
22.1829 | -35.3904 | -32.7296
22.189 | -35.3908 | -32.736
22.1951 | -35.3913 | -32.7424
22.2012 | -35.3918 | -32.7485
22.2073 | -35.3922 | -32.7547
22.2134 | -35.3927 | -32.7609
22.2195 | -35.3931 | -32.7669
22.2256 | -35.3936 | -32.7732
22.2317 | -35.394 | -32.7793
22.2378 | -35.3944 | -32.7857
22.2439 | -35.3949 | -32.7921
22.25 | -35.3953 | -32.7982
22.2561 | -35.3958 | -32.8041
22.2622 | -35.3962 | -32.8105
22.2683 | -35.3967 | -32.8162
22.2744 | -35.3972 | -32.8228
22.2806 | -35.3976 | -32.8291
22.2867 | -35.3981 | -32.835
22.2928 | -35.3985 | -32.8414
22.2989 | -35.3989 | -32.8473
22.305 | -35.3993 | -32.8532
22.3111 | -35.3998 | -32.8594
22.3172 | -35.4003 | -32.8656
22.3233 | -35.4008 | -32.8715
22.3294 | -35.4012 | -32.8778
22.3355 | -35.4017 | -32.8841
22.3416 | -35.4022 | -32.8903
22.3477 | -35.4026 | -32.8967
22.3538 | -35.4029 | -32.9027
22.3599 | -35.4034 | -32.9087
22.366 | -35.4038 | -32.9148
22.3721 | -35.4042 | -32.9207
22.3782 | -35.4046 | -32.9271
22.3843 | -35.405 | -32.9335
22.3904 | -35.4055 | -32.9397
22.3965 | -35.406 | -32.946
22.4026 | -35.4064 | -32.9521
22.4087 | -35.4068 | -32.9581
22.4148 | -35.4072 | -32.9644
22.4209 | -35.4077 | -32.9706
22.4271 | -35.4082 | -32.9764
22.4332 | -35.4086 | -32.9826
22.4393 | -35.4091 | -32.9883
22.4454 | -35.4096 | -32.9944
22.4515 | -35.41 | -33.0009
22.4576 | -35.4104 | -33.007
22.4637 | -35.4109 | -33.0129
22.4698 | -35.4114 | -33.0192
22.4759 | -35.4119 | -33.0251
22.482 | -35.4123 | -33.0315
22.4881 | -35.4128 | -33.0376
22.4942 | -35.4133 | -33.0436
22.5003 | -35.4137 | -33.0496
22.5064 | -35.4142 | -33.0563
22.5125 | -35.4147 | -33.0625
22.5186 | -35.4151 | -33.0682
22.5247 | -35.4156 | -33.0747
22.5308 | -35.4161 | -33.081
22.5369 | -35.4166 | -33.0871
22.543 | -35.417 | -33.0935
22.5491 | -35.4174 | -33.0998
22.5552 | -35.4179 | -33.1062
22.5613 | -35.4184 | -33.1126
22.5675 | -35.4189 | -33.1187
22.5736 | -35.4193 | -33.1249
22.5797 | -35.4198 | -33.1306
22.5858 | -35.4203 | -33.1366
22.5919 | -35.4208 | -33.1426
22.598 | -35.4212 | -33.1488
22.6041 | -35.4216 | -33.1552
22.6102 | -35.4221 | -33.1614
22.6163 | -35.4225 | -33.1671
22.6224 | -35.4229 | -33.1728
22.6285 | -35.4234 | -33.1791
22.6346 | -35.4238 | -33.1855
22.6407 | -35.4242 | -33.192
22.6468 | -35.4247 | -33.198
22.6529 | -35.4251 | -33.2035
22.659 | -35.4255 | -33.2097
22.6651 | -35.426 | -33.2161
22.6712 | -35.4265 | -33.2224
22.6773 | -35.427 | -33.2285
22.6834 | -35.4274 | -33.2348
22.6895 | -35.4279 | -33.2407
22.6956 | -35.4283 | -33.2467
22.7017 | -35.4287 | -33.2531
22.7079 | -35.4291 | -33.2593
22.714 | -35.4294 | -33.2653
22.7201 | -35.4299 | -33.2715
22.7262 | -35.4303 | -33.2775
22.7323 | -35.4306 | -33.2835
22.7384 | -35.431 | -33.2897
22.7445 | -35.4314 | -33.2958
22.7506 | -35.4318 | -33.302
22.7567 | -35.4323 | -33.308
22.7628 | -35.4327 | -33.3139
22.7689 | -35.4331 | -33.3201
22.775 | -35.4335 | -33.3264
22.7811 | -35.434 | -33.3328
22.7872 | -35.4344 | -33.3394
22.7933 | -35.4348 | -33.3453
22.7994 | -35.4353 | -33.3514
22.8055 | -35.4357 | -33.3579
22.8116 | -35.4362 | -33.3639
22.8177 | -35.4366 | -33.3702
22.8238 | -35.437 | -33.3766
22.8299 | -35.4374 | -33.3824
22.836 | -35.4379 | -33.3887
22.8421 | -35.4383 | -33.3949
22.8482 | -35.4388 | -33.4013
22.8544 | -35.4393 | -33.4077
22.8605 | -35.4398 | -33.4139
22.8666 | -35.4401 | -33.4205
22.8727 | -35.4405 | -33.4266
22.8788 | -35.441 | -33.4329
22.8849 | -35.4414 | -33.4384
22.891 | -35.4418 | -33.4447
22.8971 | -35.4424 | -33.451
22.9032 | -35.4428 | -33.4572
22.9093 | -35.4433 | -33.4631
22.9154 | -35.4437 | -33.4697
22.9215 | -35.4442 | -33.4757
22.9276 | -35.4446 | -33.4819
22.9337 | -35.445 | -33.4878
22.9398 | -35.4454 | -33.4936
22.9459 | -35.4459 | -33.4996
22.952 | -35.4464 | -33.5058
22.9581 | -35.4468 | -33.5119
22.9642 | -35.4472 | -33.518
22.9703 | -35.4476 | -33.5234
22.9764 | -35.4481 | -33.5296
22.9825 | -35.4486 | -33.5356
22.9886 | -35.449 | -33.5418
22.9948 | -35.4494 | -33.5477
23.0009 | -35.4498 | -33.5541
23.007 | -35.4503 | -33.5603
23.0131 | -35.4508 | -33.5665
23.0192 | -35.4512 | -33.5727
23.0253 | -35.4516 | -33.5789
23.0314 | -35.452 | -33.5851
23.0375 | -35.4525 | -33.5911
23.0436 | -35.4529 | -33.5971
23.0497 | -35.4534 | -33.6031
23.0558 | -35.4539 | -33.6091
23.0619 | -35.4543 | -33.6148
23.068 | -35.4547 | -33.6209
23.0741 | -35.455 | -33.6268
23.0802 | -35.4555 | -33.6332
23.0863 | -35.4559 | -33.6389
23.0924 | -35.4563 | -33.6447
23.0985 | -35.4568 | -33.6512
23.1046 | -35.4572 | -33.6572
23.1107 | -35.4576 | -33.6633
23.1168 | -35.4581 | -33.6692
23.1229 | -35.4585 | -33.6753
23.129 | -35.459 | -33.6815
23.1351 | -35.4594 | -33.6876
23.1413 | -35.4599 | -33.6936
23.1474 | -35.4603 | -33.6996
23.1535 | -35.4607 | -33.7056
23.1596 | -35.4611 | -33.7116
23.1657 | -35.4615 | -33.7175
23.1718 | -35.462 | -33.7235
23.1779 | -35.4624 | -33.7297
23.184 | -35.4629 | -33.7356
23.1901 | -35.4634 | -33.7418
23.1962 | -35.4638 | -33.7477
23.2023 | -35.4643 | -33.7537
23.2084 | -35.4647 | -33.7599
23.2145 | -35.4651 | -33.7661
23.2206 | -35.4656 | -33.7723
23.2267 | -35.466 | -33.7785
23.2328 | -35.4665 | -33.7844
23.2389 | -35.4668 | -33.7904
23.245 | -35.4673 | -33.7965
23.2511 | -35.4677 | -33.8028
23.2572 | -35.4681 | -33.8091
23.2633 | -35.4686 | -33.8153
23.2694 | -35.469 | -33.8217
23.2755 | -35.4695 | -33.828
23.2817 | -35.4699 | -33.8338
23.2878 | -35.4703 | -33.8398
23.2939 | -35.4707 | -33.8459
23.3 | -35.4712 | -33.8519
23.3061 | -35.4716 | -33.8581
23.3122 | -35.4721 | -33.8643
23.3183 | -35.4725 | -33.8703
23.3244 | -35.473 | -33.8763
23.3305 | -35.4734 | -33.8823
23.3366 | -35.4738 | -33.8883
23.3427 | -35.4743 | -33.8944
23.3488 | -35.4747 | -33.9003
23.3549 | -35.4751 | -33.9063
23.361 | -35.4756 | -33.9123
23.3671 | -35.476 | -33.9184
23.3732 | -35.4764 | -33.9244
23.3793 | -35.4769 | -33.9305
23.3854 | -35.4773 | -33.9366
23.3915 | -35.4777 | -33.9423
23.3976 | -35.4782 | -33.9483
23.4037 | -35.4786 | -33.9542
23.4098 | -35.479 | -33.96
23.4159 | -35.4794 | -33.966
23.422 | -35.4798 | -33.972
23.4282 | -35.4803 | -33.9776
23.4343 | -35.4808 | -33.9836
23.4404 | -35.4812 | -33.9896
23.4465 | -35.4816 | -33.9958
23.4526 | -35.4821 | -34.0019
23.4587 | -35.4825 | -34.0078
23.4648 | -35.4829 | -34.0135
23.4709 | -35.4834 | -34.0194
23.477 | -35.4838 | -34.0257
23.4831 | -35.4842 | -34.0317
23.4892 | -35.4846 | -34.0375
23.4953 | -35.4851 | -34.0436
23.5014 | -35.4855 | -34.0494
23.5075 | -35.4859 | -34.0555
23.5136 | -35.4863 | -34.0615
23.5197 | -35.4868 | -34.0673
23.5258 | -35.4872 | -34.0732
23.5319 | -35.4876 | -34.079
23.538 | -35.488 | -34.0851
23.5441 | -35.4884 | -34.0912
23.5502 | -35.4888 | -34.0971
23.5563 | -35.4892 | -34.1033
23.5624 | -35.4897 | -34.1094
23.5686 | -35.4901 | -34.1153
23.5747 | -35.4905 | -34.1215
23.5808 | -35.4909 | -34.1276
23.5869 | -35.4913 | -34.1337
23.593 | -35.4917 | -34.1396
23.5991 | -35.4922 | -34.1457
23.6052 | -35.4926 | -34.1518
23.6113 | -35.493 | -34.1579
23.6174 | -35.4935 | -34.1641
23.6235 | -35.4939 | -34.17
23.6296 | -35.4943 | -34.1761
23.6357 | -35.4948 | -34.182
23.6418 | -35.4952 | -34.1883
23.6479 | -35.4956 | -34.1941
23.654 | -35.496 | -34.2001
23.6601 | -35.4964 | -34.2064
23.6662 | -35.4968 | -34.2126
23.6723 | -35.4972 | -34.2185
23.6784 | -35.4977 | -34.2245
23.6845 | -35.4981 | -34.2305
23.6906 | -35.4985 | -34.2364
23.6967 | -35.4989 | -34.2425
23.7028 | -35.4993 | -34.2485
23.7089 | -35.4997 | -34.2543
23.7151 | -35.5001 | -34.2604
23.7212 | -35.5006 | -34.2664
23.7273 | -35.501 | -34.2723
23.7334 | -35.5014 | -34.2783
23.7395 | -35.5019 | -34.2846
23.7456 | -35.5023 | -34.2904
23.7517 | -35.5027 | -34.2964
23.7578 | -35.5031 | -34.3024
23.7639 | -35.5035 | -34.3084
23.77 | -35.504 | -34.3146
23.7761 | -35.5044 | -34.3206
23.7822 | -35.5048 | -34.3265
23.7883 | -35.5052 | -34.3326
23.7944 | -35.5056 | -34.3385
23.8005 | -35.5061 | -34.3445
23.8066 | -35.5065 | -34.3505
23.8127 | -35.5069 | -34.3565
23.8188 | -35.5073 | -34.3627
23.8249 | -35.5078 | -34.3686
23.831 | -35.5082 | -34.3745
23.8371 | -35.5086 | -34.3805
23.8432 | -35.509 | -34.3865
23.8493 | -35.5095 | -34.3923
23.8555 | -35.5099 | -34.3983
23.8616 | -35.5103 | -34.4043
23.8677 | -35.5107 | -34.4104
23.8738 | -35.5111 | -34.4163
23.8799 | -35.5116 | -34.4225
23.886 | -35.512 | -34.4283
23.8921 | -35.5124 | -34.4341
23.8982 | -35.5128 | -34.44
23.9043 | -35.5132 | -34.4461
23.9104 | -35.5137 | -34.4521
23.9165 | -35.5141 | -34.4582
23.9226 | -35.5145 | -34.4642
23.9287 | -35.5149 | -34.4702
23.9348 | -35.5153 | -34.4761
23.9409 | -35.5157 | -34.4821
23.947 | -35.5161 | -34.4879
23.9531 | -35.5165 | -34.4937
23.9592 | -35.5169 | -34.4999
23.9653 | -35.5173 | -34.506
23.9714 | -35.5177 | -34.5119
23.9775 | -35.5181 | -34.518
23.9836 | -35.5186 | -34.5241
23.9897 | -35.519 | -34.5299
23.9958 | -35.5194 | -34.5357
24.002 | -35.5198 | -34.5416
24.0081 | -35.5203 | -34.5475
24.0142 | -35.5207 | -34.5534
24.0203 | -35.5211 | -34.5593
24.0264 | -35.5215 | -34.5652
24.0325 | -35.5219 | -34.5711
24.0386 | -35.5223 | -34.5772
24.0447 | -35.5227 | -34.5832
24.0508 | -35.5232 | -34.589
24.0569 | -35.5236 | -34.595
24.063 | -35.524 | -34.6007
24.0691 | -35.5244 | -34.6067
24.0752 | -35.5248 | -34.6129
24.0813 | -35.5253 | -34.6189
24.0874 | -35.5257 | -34.6247
24.0935 | -35.5261 | -34.6309
24.0996 | -35.5265 | -34.6366
24.1057 | -35.527 | -34.6426
24.1118 | -35.5274 | -34.6486
24.1179 | -35.5278 | -34.6544
24.124 | -35.5282 | -34.6604
24.1301 | -35.5287 | -34.6661
24.1362 | -35.5291 | -34.6719
24.1424 | -35.5295 | -34.6778
24.1485 | -35.5299 | -34.6841
24.1546 | -35.5304 | -34.6899
24.1607 | -35.5308 | -34.6962
24.1668 | -35.5312 | -34.7019
24.1729 | -35.5316 | -34.7079
24.179 | -35.532 | -34.7139
24.1851 | -35.5324 | -34.7198
24.1912 | -35.5328 | -34.7257
24.1973 | -35.5332 | -34.7316
24.2034 | -35.5337 | -34.7375
24.2095 | -35.5341 | -34.7435
24.2156 | -35.5345 | -34.7494
24.2217 | -35.5349 | -34.7552
24.2278 | -35.5354 | -34.7611
24.2339 | -35.5358 | -34.7668
24.24 | -35.5362 | -34.7728
24.2461 | -35.5366 | -34.7787
24.2522 | -35.537 | -34.7848
24.2583 | -35.5374 | -34.7908
24.2644 | -35.5379 | -34.797
24.2705 | -35.5383 | -34.8028
24.2766 | -35.5387 | -34.8085
24.2827 | -35.5391 | -34.8144
24.2889 | -35.5395 | -34.8203
24.295 | -35.5399 | -34.8262
24.3011 | -35.5403 | -34.8323
24.3072 | -35.5407 | -34.8382
24.3133 | -35.5411 | -34.844
24.3194 | -35.5415 | -34.8499
24.3255 | -35.542 | -34.8561
24.3316 | -35.5423 | -34.8621
24.3377 | -35.5428 | -34.868
24.3438 | -35.5432 | -34.8737
24.3499 | -35.5436 | -34.8795
24.356 | -35.5439 | -34.8856
24.3621 | -35.5443 | -34.8913
24.3682 | -35.5447 | -34.8973
24.3743 | -35.5452 | -34.9032
24.3804 | -35.5456 | -34.9089
24.3865 | -35.546 | -34.9149
24.3926 | -35.5464 | -34.9208
24.3987 | -35.5468 | -34.9269
24.4048 | -35.5472 | -34.9328
24.4109 | -35.5477 | -34.9385
24.417 | -35.5481 | -34.9444
24.4231 | -35.5485 | -34.9502
24.4293 | -35.5489 | -34.9561
24.4354 | -35.5493 | -34.962
24.4415 | -35.5498 | -34.9679
24.4476 | -35.5502 | -34.9736
24.4537 | -35.5506 | -34.9795
24.4598 | -35.551 | -34.9854
24.4659 | -35.5514 | -34.9914
24.472 | -35.5518 | -34.9974
24.4781 | -35.5522 | -35.0034
24.4842 | -35.5526 | -35.0092
24.4903 | -35.553 | -35.0151
24.4964 | -35.5535 | -35.0211
24.5025 | -35.5539 | -35.0271
24.5086 | -35.5543 | -35.0332
24.5147 | -35.5547 | -35.0392
24.5208 | -35.5551 | -35.0449
24.5269 | -35.5555 | -35.0508
24.533 | -35.5559 | -35.0567
24.5391 | -35.5564 | -35.0625
24.5452 | -35.5568 | -35.0684
24.5513 | -35.5572 | -35.074
24.5574 | -35.5576 | -35.0799
24.5635 | -35.558 | -35.0858
24.5696 | -35.5584 | -35.0918
24.5758 | -35.5589 | -35.0975
24.5819 | -35.5593 | -35.1034
24.588 | -35.5597 | -35.1092
24.5941 | -35.5601 | -35.1151
24.6002 | -35.5605 | -35.121
24.6063 | -35.5609 | -35.1269
24.6124 | -35.5613 | -35.1328
24.6185 | -35.5618 | -35.1387
24.6246 | -35.5621 | -35.1444
24.6307 | -35.5625 | -35.1503
24.6368 | -35.5629 | -35.1562
24.6429 | -35.5634 | -35.1623
24.649 | -35.5638 | -35.1683
24.6551 | -35.5642 | -35.1742
24.6612 | -35.5646 | -35.1798
24.6673 | -35.565 | -35.1856
24.6734 | -35.5654 | -35.1915
24.6795 | -35.5658 | -35.1974
24.6856 | -35.5662 | -35.2034
24.6917 | -35.5667 | -35.2091
24.6978 | -35.5671 | -35.2149
24.7039 | -35.5675 | -35.2208
24.71 | -35.5679 | -35.2268
24.7162 | -35.5683 | -35.2328
24.7223 | -35.5687 | -35.2387
24.7284 | -35.5691 | -35.2444
24.7345 | -35.5695 | -35.2502
24.7406 | -35.5699 | -35.2561
24.7467 | -35.5703 | -35.2619
24.7528 | -35.5707 | -35.2677
24.7589 | -35.5711 | -35.2736
24.765 | -35.5715 | -35.2793
24.7711 | -35.5719 | -35.2851
24.7772 | -35.5723 | -35.2909
24.7833 | -35.5727 | -35.2966
24.7894 | -35.5731 | -35.3025
24.7955 | -35.5735 | -35.3084
24.8016 | -35.574 | -35.3141
24.8077 | -35.5743 | -35.32
24.8138 | -35.5748 | -35.3261
24.8199 | -35.5752 | -35.3318
24.826 | -35.5756 | -35.3376
24.8321 | -35.576 | -35.3435
24.8382 | -35.5764 | -35.3492
24.8443 | -35.5768 | -35.355
24.8504 | -35.5772 | -35.3607
24.8565 | -35.5776 | -35.3665
24.8627 | -35.578 | -35.3723
24.8688 | -35.5785 | -35.378
24.8749 | -35.5789 | -35.3839
24.881 | -35.5793 | -35.3899
24.8871 | -35.5797 | -35.3958
24.8932 | -35.5801 | -35.4016
24.8993 | -35.5805 | -35.4073
24.9054 | -35.5809 | -35.413
24.9115 | -35.5813 | -35.4189
24.9176 | -35.5817 | -35.4248
24.9237 | -35.5821 | -35.4306
24.9298 | -35.5825 | -35.4364
24.9359 | -35.5829 | -35.4422
24.942 | -35.5833 | -35.4477
24.9481 | -35.5837 | -35.4536
24.9542 | -35.5841 | -35.4594
24.9603 | -35.5845 | -35.4653
24.9664 | -35.5849 | -35.4712
24.9725 | -35.5853 | -35.4772
24.9786 | -35.5857 | -35.4828
24.9847 | -35.5861 | -35.4887
24.9908 | -35.5865 | -35.4944
24.9969 | -35.5869 | -35.5002
25.0031 | -35.5873 | -35.5059
25.0092 | -35.5877 | -35.5115
25.0153 | -35.588 | -35.5172
25.0214 | -35.5885 | -35.5231
25.0275 | -35.5888 | -35.5288
25.0336 | -35.5893 | -35.5347
25.0397 | -35.5897 | -35.5405
25.0458 | -35.5901 | -35.5463
25.0519 | -35.5905 | -35.5522
25.058 | -35.5909 | -35.5579
25.0641 | -35.5913 | -35.5637
25.0702 | -35.5917 | -35.5696
25.0763 | -35.5921 | -35.5755
25.0824 | -35.5925 | -35.581
25.0885 | -35.5929 | -35.5868
25.0946 | -35.5933 | -35.5925
25.1007 | -35.5937 | -35.5983
25.1068 | -35.5941 | -35.604
25.1129 | -35.5945 | -35.61
25.119 | -35.5949 | -35.6156
25.1251 | -35.5953 | -35.6213
25.1312 | -35.5957 | -35.6272
25.1373 | -35.5961 | -35.6331
25.1435 | -35.5965 | -35.639
25.1496 | -35.5969 | -35.6446
25.1557 | -35.5973 | -35.6505
25.1618 | -35.5977 | -35.6564
25.1679 | -35.5981 | -35.6622
25.174 | -35.5985 | -35.6679
25.1801 | -35.5989 | -35.6735
25.1862 | -35.5993 | -35.6792
25.1923 | -35.5997 | -35.685
25.1984 | -35.6001 | -35.6906
25.2045 | -35.6005 | -35.6965
25.2106 | -35.6009 | -35.7024
25.2167 | -35.6013 | -35.7082
25.2228 | -35.6017 | -35.7139
25.2289 | -35.6022 | -35.7196
25.235 | -35.6026 | -35.7254
25.2411 | -35.603 | -35.7311
25.2472 | -35.6034 | -35.7369
25.2533 | -35.6038 | -35.7426
25.2594 | -35.6042 | -35.7479
25.2655 | -35.6046 | -35.7537
25.2716 | -35.605 | -35.7594
25.2777 | -35.6054 | -35.765
25.2838 | -35.6058 | -35.7708
25.29 | -35.6062 | -35.7766
25.2961 | -35.6066 | -35.7823
25.3022 | -35.607 | -35.7879
25.3083 | -35.6074 | -35.7936
25.3144 | -35.6078 | -35.7992
25.3205 | -35.6081 | -35.805
25.3266 | -35.6085 | -35.8106
25.3327 | -35.6089 | -35.8165
25.3388 | -35.6093 | -35.8223
25.3449 | -35.6097 | -35.8279
25.351 | -35.6101 | -35.8336
25.3571 | -35.6105 | -35.8394
25.3632 | -35.6109 | -35.845
25.3693 | -35.6113 | -35.8507
25.3754 | -35.6117 | -35.8564
25.3815 | -35.6121 | -35.8622
25.3876 | -35.6125 | -35.8679
25.3937 | -35.6129 | -35.8736
25.3998 | -35.6133 | -35.879
25.4059 | -35.6137 | -35.8847
25.412 | -35.6141 | -35.8904
25.4181 | -35.6145 | -35.8961
25.4242 | -35.6149 | -35.9019
25.4304 | -35.6153 | -35.9074
25.4365 | -35.6157 | -35.913
25.4426 | -35.616 | -35.9187
25.4487 | -35.6164 | -35.9245
25.4548 | -35.6168 | -35.9302
25.4609 | -35.6172 | -35.936
25.467 | -35.6176 | -35.9417
25.4731 | -35.618 | -35.9474
25.4792 | -35.6185 | -35.9532
25.4853 | -35.6189 | -35.9589
25.4914 | -35.6193 | -35.9647
25.4975 | -35.6197 | -35.9703
25.5036 | -35.6201 | -35.9759
25.5097 | -35.6205 | -35.9815
25.5158 | -35.6209 | -35.9872
25.5219 | -35.6213 | -35.9929
25.528 | -35.6217 | -35.9986
25.5341 | -35.6221 | -36.0042
25.5402 | -35.6225 | -36.0099
25.5463 | -35.6229 | -36.0155
25.5524 | -35.6233 | -36.0212
25.5585 | -35.6237 | -36.0268
25.5646 | -35.6241 | -36.0327
25.5707 | -35.6245 | -36.0385
25.5769 | -35.6249 | -36.0441
25.583 | -35.6253 | -36.0498
25.5891 | -35.6257 | -36.0555
25.5952 | -35.6261 | -36.0614
25.6013 | -35.6265 | -36.0669
25.6074 | -35.6269 | -36.0725
25.6135 | -35.6273 | -36.0782
25.6196 | -35.6277 | -36.0836
25.6257 | -35.6281 | -36.0894
25.6318 | -35.6285 | -36.095
25.6379 | -35.6289 | -36.1007
25.644 | -35.6292 | -36.1061
25.6501 | -35.6296 | -36.1117
25.6562 | -35.63 | -36.1173
25.6623 | -35.6304 | -36.123
25.6684 | -35.6308 | -36.1288
25.6745 | -35.6312 | -36.1345
25.6806 | -35.6316 | -36.1398
25.6867 | -35.6319 | -36.1455
25.6928 | -35.6324 | -36.1513
25.6989 | -35.6327 | -36.1568
25.705 | -35.6331 | -36.1624
25.7111 | -35.6335 | -36.1682
25.7173 | -35.6339 | -36.1736
25.7234 | -35.6343 | -36.1793
25.7295 | -35.6346 | -36.185
25.7356 | -35.6351 | -36.1906
25.7417 | -35.6354 | -36.1962
25.7478 | -35.6359 | -36.2016
25.7539 | -35.6362 | -36.2071
25.76 | -35.6366 | -36.2127
25.7661 | -35.637 | -36.2181
25.7722 | -35.6374 | -36.224
25.7783 | -35.6378 | -36.2295
25.7844 | -35.6382 | -36.2351
25.7905 | -35.6386 | -36.2407
25.7966 | -35.639 | -36.2463
25.8027 | -35.6393 | -36.2517
25.8088 | -35.6397 | -36.2574
25.8149 | -35.6401 | -36.2629
25.821 | -35.6405 | -36.2683
25.8271 | -35.6409 | -36.274
25.8332 | -35.6413 | -36.2796
25.8393 | -35.6417 | -36.2852
25.8454 | -35.642 | -36.2908
25.8515 | -35.6424 | -36.2965
25.8576 | -35.6428 | -36.3018
25.8638 | -35.6432 | -36.3075
25.8699 | -35.6436 | -36.3129
25.876 | -35.6439 | -36.3185
25.8821 | -35.6443 | -36.3241
25.8882 | -35.6447 | -36.3294
25.8943 | -35.6451 | -36.3351
25.9004 | -35.6454 | -36.3407
25.9065 | -35.6458 | -36.3464
25.9126 | -35.6462 | -36.352
25.9187 | -35.6466 | -36.3575
25.9248 | -35.647 | -36.363
25.9309 | -35.6474 | -36.3685
25.937 | -35.6478 | -36.3741
25.9431 | -35.6482 | -36.3797
25.9492 | -35.6485 | -36.3853
25.9553 | -35.6489 | -36.3909
25.9614 | -35.6493 | -36.3964
25.9675 | -35.6498 | -36.402
25.9736 | -35.6502 | -36.4077
25.9797 | -35.6505 | -36.4134
25.9858 | -35.6509 | -36.4189
25.9919 | -35.6513 | -36.4246
25.998 | -35.6517 | -36.4298
26.0042 | -35.6521 | -36.4354
26.0103 | -35.6525 | -36.4409
26.0164 | -35.6529 | -36.4464
26.0225 | -35.6532 | -36.4519
26.0286 | -35.6536 | -36.4574
26.0347 | -35.654 | -36.4627
26.0408 | -35.6544 | -36.4682
26.0469 | -35.6548 | -36.4738
26.053 | -35.6551 | -36.4794
26.0591 | -35.6555 | -36.485
26.0652 | -35.6559 | -36.4904
26.0713 | -35.6562 | -36.4959
26.0774 | -35.6566 | -36.5015
26.0835 | -35.657 | -36.5071
26.0896 | -35.6574 | -36.5125
26.0957 | -35.6578 | -36.5182
26.1018 | -35.6582 | -36.5233
26.1079 | -35.6586 | -36.5287
26.114 | -35.659 | -36.5341
26.1201 | -35.6593 | -36.5399
26.1262 | -35.6597 | -36.5454
26.1323 | -35.6601 | -36.551
26.1384 | -35.6605 | -36.5565
26.1445 | -35.6609 | -36.5622
26.1507 | -35.6613 | -36.568
26.1568 | -35.6617 | -36.5734
26.1629 | -35.6621 | -36.579
26.169 | -35.6625 | -36.5845
26.1751 | -35.6629 | -36.5897
26.1812 | -35.6633 | -36.5954
26.1873 | -35.6636 | -36.6009
26.1934 | -35.664 | -36.6061
26.1995 | -35.6644 | -36.6117
26.2056 | -35.6647 | -36.6171
26.2117 | -35.6652 | -36.6225
26.2178 | -35.6656 | -36.628
26.2239 | -35.6659 | -36.6333
26.23 | -35.6663 | -36.6387
26.2361 | -35.6667 | -36.6444
26.2422 | -35.6671 | -36.6496
26.2483 | -35.6675 | -36.655
26.2544 | -35.6678 | -36.6604
26.2605 | -35.6682 | -36.6659
26.2666 | -35.6686 | -36.6715
26.2727 | -35.669 | -36.6769
26.2788 | -35.6694 | -36.682
26.2849 | -35.6698 | -36.6873
26.2911 | -35.6701 | -36.6928
26.2972 | -35.6705 | -36.6983
26.3033 | -35.6709 | -36.7037
26.3094 | -35.6713 | -36.7092
26.3155 | -35.6717 | -36.7146
26.3216 | -35.6721 | -36.7202
26.3277 | -35.6724 | -36.7257
26.3338 | -35.6728 | -36.7312
26.3399 | -35.6732 | -36.7368
26.346 | -35.6736 | -36.7423
26.3521 | -35.674 | -36.7478
26.3582 | -35.6743 | -36.7533
26.3643 | -35.6747 | -36.7588
26.3704 | -35.6751 | -36.7642
26.3765 | -35.6755 | -36.7698
26.3826 | -35.6759 | -36.7748
26.3887 | -35.6763 | -36.7804
26.3948 | -35.6766 | -36.7859
26.4009 | -35.677 | -36.7914
26.407 | -35.6774 | -36.797
26.4131 | -35.6778 | -36.8026
26.4192 | -35.6782 | -36.8076
26.4253 | -35.6785 | -36.813
26.4314 | -35.6789 | -36.8183
26.4376 | -35.6793 | -36.8239
26.4437 | -35.6797 | -36.8292
26.4498 | -35.68 | -36.8346
26.4559 | -35.6804 | -36.8395
26.462 | -35.6808 | -36.8451
26.4681 | -35.6811 | -36.8505
26.4742 | -35.6815 | -36.856
26.4803 | -35.6819 | -36.8614
26.4864 | -35.6822 | -36.8668
26.4925 | -35.6826 | -36.8722
26.4986 | -35.683 | -36.8776
26.5047 | -35.6834 | -36.883
26.5108 | -35.6837 | -36.8883
26.5169 | -35.6841 | -36.8939
26.523 | -35.6845 | -36.8988
26.5291 | -35.6848 | -36.9041
26.5352 | -35.6852 | -36.9095
26.5413 | -35.6856 | -36.9148
26.5474 | -35.6859 | -36.9203
26.5535 | -35.6863 | -36.9259
26.5596 | -35.6867 | -36.9313
26.5657 | -35.6871 | -36.9367
26.5718 | -35.6874 | -36.9422
26.578 | -35.6878 | -36.9476
26.5841 | -35.6882 | -36.9531
26.5902 | -35.6885 | -36.9585
26.5963 | -35.6889 | -36.9636
26.6024 | -35.6893 | -36.9688
26.6085 | -35.6896 | -36.9743
26.6146 | -35.69 | -36.9797
26.6207 | -35.6904 | -36.985
26.6268 | -35.6908 | -36.9905
26.6329 | -35.6911 | -36.9958
26.639 | -35.6915 | -37.0011
26.6451 | -35.6919 | -37.0066
26.6512 | -35.6923 | -37.0118
26.6573 | -35.6927 | -37.0172
26.6634 | -35.6931 | -37.0223
26.6695 | -35.6934 | -37.0276
26.6756 | -35.6938 | -37.0331
26.6817 | -35.6941 | -37.0384
26.6878 | -35.6945 | -37.0437
26.6939 | -35.6949 | -37.0491
26.7 | -35.6953 | -37.0541
26.7061 | -35.6956 | -37.0595
26.7122 | -35.696 | -37.0648
26.7183 | -35.6964 | -37.07
26.7245 | -35.6968 | -37.0752
26.7306 | -35.6972 | -37.0804
26.7367 | -35.6975 | -37.0857
26.7428 | -35.6979 | -37.091
26.7489 | -35.6983 | -37.0966
26.755 | -35.6987 | -37.1018
26.7611 | -35.699 | -37.107
26.7672 | -35.6994 | -37.1122
26.7733 | -35.6998 | -37.1172
26.7794 | -35.7002 | -37.1225
26.7855 | -35.7005 | -37.1278
26.7916 | -35.7009 | -37.1331
26.7977 | -35.7013 | -37.1383
26.8038 | -35.7017 | -37.1433
26.8099 | -35.702 | -37.1486
26.816 | -35.7024 | -37.1541
26.8221 | -35.7027 | -37.1593
26.8282 | -35.7031 | -37.1646
26.8343 | -35.7034 | -37.1698
26.8404 | -35.7038 | -37.175
26.8465 | -35.7042 | -37.1805
26.8526 | -35.7046 | -37.1859
26.8587 | -35.7049 | -37.1911
26.8649 | -35.7053 | -37.1965
26.871 | -35.7056 | -37.2018
26.8771 | -35.7059 | -37.207
26.8832 | -35.7063 | -37.2126
26.8893 | -35.7066 | -37.2178
26.8954 | -35.707 | -37.2231
26.9015 | -35.7073 | -37.2287
26.9076 | -35.7077 | -37.2338
26.9137 | -35.708 | -37.2385
26.9198 | -35.7084 | -37.2441
26.9259 | -35.7088 | -37.2493
26.932 | -35.7091 | -37.2546
26.9381 | -35.7095 | -37.2598
26.9442 | -35.7099 | -37.2651
26.9503 | -35.7102 | -37.2703
26.9564 | -35.7106 | -37.2757
26.9625 | -35.711 | -37.2812
26.9686 | -35.7114 | -37.2864
26.9747 | -35.7117 | -37.2918
26.9808 | -35.7121 | -37.2968
26.9869 | -35.7125 | -37.3021
26.993 | -35.7129 | -37.3074
26.9991 | -35.7132 | -37.3129
27.0052 | -35.7136 | -37.3183
27.0114 | -35.7139 | -37.3235
27.0175 | -35.7143 | -37.3285
27.0236 | -35.7146 | -37.3338
27.0297 | -35.715 | -37.339
27.0358 | -35.7154 | -37.3443
27.0419 | -35.7158 | -37.3496
27.048 | -35.7161 | -37.355
27.0541 | -35.7165 | -37.3602
27.0602 | -35.7169 | -37.3653
27.0663 | -35.7172 | -37.3708
27.0724 | -35.7175 | -37.3759
27.0785 | -35.7179 | -37.3813
27.0846 | -35.7183 | -37.3864
27.0907 | -35.7186 | -37.3918
27.0968 | -35.719 | -37.3971
27.1029 | -35.7193 | -37.4023
27.109 | -35.7197 | -37.4077
27.1151 | -35.72 | -37.4131
27.1212 | -35.7204 | -37.4179
27.1273 | -35.7208 | -37.4231
27.1334 | -35.7212 | -37.4282
27.1395 | -35.7215 | -37.4335
27.1456 | -35.7219 | -37.4385
27.1518 | -35.7223 | -37.4438
27.1579 | -35.7227 | -37.4487
27.164 | -35.723 | -37.4539
27.1701 | -35.7234 | -37.459
27.1762 | -35.7238 | -37.464
27.1823 | -35.7241 | -37.469
27.1884 | -35.7245 | -37.4742
27.1945 | -35.7249 | -37.4793
27.2006 | -35.7252 | -37.4844
27.2067 | -35.7256 | -37.4896
27.2128 | -35.726 | -37.4948
27.2189 | -35.7263 | -37.4999
27.225 | -35.7267 | -37.5052
27.2311 | -35.727 | -37.5102
27.2372 | -35.7274 | -37.5152
27.2433 | -35.7277 | -37.5203
27.2494 | -35.7281 | -37.5255
27.2555 | -35.7285 | -37.5307
27.2616 | -35.7288 | -37.5354
27.2677 | -35.7292 | -37.5408
27.2738 | -35.7295 | -37.5458
27.2799 | -35.7299 | -37.5509
27.286 | -35.7303 | -37.5561
27.2921 | -35.7307 | -37.5614
27.2983 | -35.731 | -37.566
27.3044 | -35.7313 | -37.5714
27.3105 | -35.7316 | -37.5766
27.3166 | -35.732 | -37.5818
27.3227 | -35.7324 | -37.5869
27.3288 | -35.7328 | -37.5922
27.3349 | -35.7331 | -37.597
27.341 | -35.7334 | -37.6023
27.3471 | -35.7338 | -37.6073
27.3532 | -35.7341 | -37.6125
27.3593 | -35.7345 | -37.6177
27.3654 | -35.7349 | -37.623
27.3715 | -35.7352 | -37.6278
27.3776 | -35.7356 | -37.6332
27.3837 | -35.7359 | -37.6384
27.3898 | -35.7363 | -37.6433
27.3959 | -35.7366 | -37.6483
27.402 | -35.737 | -37.6532
27.4081 | -35.7373 | -37.6584
27.4142 | -35.7377 | -37.6636
27.4203 | -35.7381 | -37.6688
27.4264 | -35.7384 | -37.6737
27.4325 | -35.7388 | -37.6788
27.4387 | -35.7392 | -37.6838
27.4448 | -35.7395 | -37.6889
27.4509 | -35.7398 | -37.6941
27.457 | -35.7402 | -37.6992
27.4631 | -35.7405 | -37.7045
27.4692 | -35.7409 | -37.7096
27.4753 | -35.7412 | -37.7145
27.4814 | -35.7415 | -37.7196
27.4875 | -35.7419 | -37.7247
27.4936 | -35.7423 | -37.7299
27.4997 | -35.7426 | -37.7352
27.5058 | -35.743 | -37.7405
27.5119 | -35.7433 | -37.7455
27.518 | -35.7437 | -37.7507
27.5241 | -35.744 | -37.7558
27.5302 | -35.7444 | -37.7609
27.5363 | -35.7448 | -37.7662
27.5424 | -35.7451 | -37.7711
27.5485 | -35.7454 | -37.7764
27.5546 | -35.7458 | -37.7815
27.5607 | -35.7462 | -37.7867
27.5668 | -35.7465 | -37.7917
27.5729 | -35.7469 | -37.7966
27.5791 | -35.7472 | -37.8013
27.5852 | -35.7476 | -37.8064
27.5913 | -35.7479 | -37.8115
27.5974 | -35.7483 | -37.8167
27.6035 | -35.7486 | -37.8219
27.6096 | -35.7489 | -37.827
27.6157 | -35.7493 | -37.832
27.6218 | -35.7496 | -37.8371
27.6279 | -35.75 | -37.8421
27.634 | -35.7504 | -37.8472
27.6401 | -35.7507 | -37.8523
27.6462 | -35.7511 | -37.8574
27.6523 | -35.7514 | -37.8622
27.6584 | -35.7518 | -37.8673
27.6645 | -35.7521 | -37.8725
27.6706 | -35.7524 | -37.8777
27.6767 | -35.7528 | -37.8827
27.6828 | -35.7531 | -37.8877
27.6889 | -35.7535 | -37.8928
27.695 | -35.7538 | -37.8979
27.7011 | -35.7541 | -37.9029
27.7072 | -35.7545 | -37.908
27.7133 | -35.7548 | -37.9131
27.7194 | -35.7552 | -37.9181
27.7256 | -35.7555 | -37.9233
27.7317 | -35.7558 | -37.9282
27.7378 | -35.7562 | -37.9332
27.7439 | -35.7566 | -37.9384
27.75 | -35.7569 | -37.9435
27.7561 | -35.7572 | -37.9484
27.7622 | -35.7576 | -37.9536
27.7683 | -35.7579 | -37.9586
27.7744 | -35.7583 | -37.9637
27.7805 | -35.7586 | -37.9687
27.7866 | -35.759 | -37.9738
27.7927 | -35.7593 | -37.9786
27.7988 | -35.7597 | -37.9837
27.8049 | -35.7601 | -37.9887
27.811 | -35.7604 | -37.9937
27.8171 | -35.7608 | -37.9987
27.8232 | -35.7611 | -38.0035
27.8293 | -35.7615 | -38.0084
27.8354 | -35.7618 | -38.0135
27.8415 | -35.7621 | -38.0186
27.8476 | -35.7625 | -38.0236
27.8537 | -35.7628 | -38.0286
27.8598 | -35.7631 | -38.0335
27.866 | -35.7635 | -38.0384
27.8721 | -35.7638 | -38.0434
27.8782 | -35.7642 | -38.0486
27.8843 | -35.7645 | -38.0536
27.8904 | -35.7649 | -38.0585
27.8965 | -35.7652 | -38.0633
27.9026 | -35.7655 | -38.0684
27.9087 | -35.7659 | -38.0734
27.9148 | -35.7662 | -38.0784
27.9209 | -35.7665 | -38.0834
27.927 | -35.7668 | -38.0883
27.9331 | -35.7671 | -38.0932
27.9392 | -35.7675 | -38.0983
27.9453 | -35.7678 | -38.1033
27.9514 | -35.7682 | -38.1084
27.9575 | -35.7685 | -38.1134
27.9636 | -35.7688 | -38.1183
27.9697 | -35.7691 | -38.123
27.9758 | -35.7695 | -38.1281
27.9819 | -35.7698 | -38.1332
27.988 | -35.7701 | -38.138
27.9941 | -35.7705 | -38.143
28.0002 | -35.7708 | -38.1477
28.0063 | -35.7711 | -38.1527
28.0125 | -35.7714 | -38.1577
28.0186 | -35.7717 | -38.1628
28.0247 | -35.772 | -38.1678
28.0308 | -35.7724 | -38.1728
28.0369 | -35.7727 | -38.1776
28.043 | -35.773 | -38.1827
28.0491 | -35.7734 | -38.1876
28.0552 | -35.7737 | -38.1928
28.0613 | -35.774 | -38.1978
28.0674 | -35.7744 | -38.203
28.0735 | -35.7747 | -38.2078

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.503 | -43.2325
21.4908 | -35.5477 | -35.2591
16.4948 | -34.9234 | -27.895
12.6602 | -34.5542 | -21.6472
9.7171 | -34.3303 | -16.7444
7.4582 | -34.1713 | -12.7611
5.7244 | -34.0578 | -9.5325
4.3936 | -33.9874 | -6.9225
3.3722 | -33.9651 | -4.8618
2.5883 | -33.9547 | -3.2957
1.9866 | -33.9484 | -2.0401
1.5248 | -33.9484 | -1.0383
1.1703 | -33.9495 | -0.37949
1 | -33.9579 | -0.053109
0.89822 | -33.9498 | 0.21644
0.68941 | -33.9412 | 0.80135
0.52919 | -33.9296 | 1.2574
0.40612 | -33.9303 | 1.8705
0.31171 | -33.9368 | 2.3235
0.23928 | -33.9378 | 3.0507
0.18358 | -33.9359 | 3.4751
0.14091 | -33.9502 | 4.3513
0.10819 | -33.9586 | 5.5107
0.083017 | -34.0048 | 6.9738
0.063705 | -34.0383 | 8.8452
0.048923 | -34.1116 | 11.331
0.037479 | -34.2171 | 14.5078
0.028801 | -34.3983 | 18.521
0.022078 | -34.6979 | 23.7959
0.017023 | -35.1701 | 30.1655
0.013018 | -35.9574 | 37.84
0.010014 | -37.0621 | 45.7344

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003704 | -12.2968
0.007407 | -30.3503
0.011111 | -37.2723
0.014815 | -42.4468
0.018519 | -45.4295
0.022222 | -47.2077
0.025926 | -48.3825
0.02963 | -49.6294
0.033333 | -51.7273
0.037037 | -53.1385
0.040741 | -52.9239
0.044444 | -53.2584
0.048148 | -54.7754
0.051852 | -56.4879
0.055556 | -58.3138
0.059259 | -58.8786
0.062963 | -57.9194
0.066667 | -58.3934
0.07037 | -60.1198
0.074074 | -60.6828
0.077778 | -61.5259
0.081481 | -62.1195
0.085185 | -63.102
0.088889 | -62.9697
0.092593 | -62.9155
0.096296 | -63.409
0.1 | -63.1551
0.1037 | -63.5986
0.10741 | -64.4933
0.11111 | -64.1555
0.11482 | -65.1573
0.11852 | -65.3173
0.12222 | -65.1121
0.12593 | -65.3637
0.12963 | -67.4455
0.13333 | -68.7062
0.13704 | -68.762
0.14074 | -68.2947
0.14444 | -67.1501
0.14815 | -67.5153
0.15185 | -68.4159
0.15556 | -67.7196
0.15926 | -67.95
0.16296 | -68.9492
0.16667 | -69.295
0.17037 | -69.7545
0.17407 | -68.9632
0.17778 | -69.2675
0.18148 | -69.911
0.18518 | -69.6992
0.18889 | -69.9188
0.19259 | -70.6892
0.1963 | -70.0091
0.2 | -70.0558
0.2037 | -71.3135
0.20741 | -72.1831
0.21111 | -72.7405
0.21482 | -71.4835
0.21852 | -70.9722
0.22222 | -71.7779
0.22593 | -71.4279
0.22963 | -72.2309
0.23333 | -72.9164
0.23704 | -72.7049
0.24074 | -71.8071
0.24444 | -72.6223
0.24815 | -72.9275
0.25185 | -73.2514
0.25556 | -73.5579
0.25926 | -73.2782
0.26296 | -72.8676
0.26667 | -73.2152
0.27037 | -73.363
0.27407 | -73.9741
0.27778 | -74.1267
0.28148 | -74.2953
0.28519 | -74.6283
0.28889 | -74.8864
0.29259 | -74.9195
0.2963 | -74.8417
0.3 | -74.1796
0.3037 | -75.6585
0.30741 | -75.9825
0.31111 | -76.311
0.31482 | -74.6084
0.31852 | -75.6096
0.32222 | -76.3299
0.32593 | -76.0191
0.32963 | -76.3967
0.33333 | -76.6301
0.33704 | -76.1591
0.34074 | -77.0519
0.34444 | -76.462
0.34815 | -75.8584
0.35185 | -76.1983
0.35556 | -76.4343
0.35926 | -77.3861
0.36296 | -77.1037
0.36667 | -76.2422
0.37037 | -76.9737
0.37407 | -76.8331
0.37778 | -75.744
0.38148 | -76.0478
0.38518 | -76.7232
0.38889 | -77.0738
0.39259 | -77.5171
0.3963 | -77.2991
0.4 | -76.4168
0.4037 | -76.287
0.40741 | -76.9564
0.41111 | -77.2545
0.41481 | -78.0749
0.41852 | -77.9458
0.42222 | -77.2759
0.42593 | -77.8454
0.42963 | -77.8484
0.43333 | -78.7631
0.43704 | -79.603
0.44074 | -79.3942
0.44444 | -78.4682
0.44815 | -78.3215
0.45185 | -78.5776
0.45556 | -79.3849
0.45926 | -77.6201
0.46296 | -78.3682
0.46667 | -79.3216
0.47037 | -79.1162
0.47407 | -80.014
0.47778 | -79.4638
0.48148 | -78.5097
0.48518 | -79.0759
0.48889 | -79.7448
0.49259 | -79.5386
0.4963 | -79.61
0.5 | -79.3965
0.5037 | -79.8474
0.50741 | -80.4057
0.51111 | -79.9929
0.51482 | -79.1953
0.51852 | -78.4818
0.52222 | -79.0385
0.52593 | -79.8469
0.52963 | -79.4688
0.53333 | -79.8596
0.53704 | -79.7036
0.54074 | -79.8699
0.54444 | -80.5404
0.54815 | -80.9618
0.55185 | -80.9317
0.55556 | -80.9105
0.55926 | -81.0068
0.56296 | -80.9619
0.56667 | -80.8089
0.57037 | -80.4793
0.57407 | -80.3903
0.57778 | -80.6857
0.58148 | -80.8947
0.58518 | -80.8286
0.58889 | -80.7591
0.59259 | -80.9107
0.5963 | -81.3352
0.6 | -81.7572
0.6037 | -82.0793
0.60741 | -81.9138
0.61111 | -81.6961
0.61482 | -81.6791
0.61852 | -81.3477
0.62222 | -80.9692
0.62593 | -81.3347
0.62963 | -82.2809
0.63333 | -82.7548
0.63704 | -82.5171
0.64074 | -82.3852
0.64444 | -82.7459
0.64815 | -82.802
0.65185 | -82.5735
0.65556 | -82.5088
0.65926 | -82.9434
0.66296 | -83.4152
0.66667 | -83.0614
0.67037 | -82.2819
0.67407 | -81.8447
0.67778 | -81.7092
0.68148 | -81.787
0.68519 | -81.9298
0.68889 | -81.9356
0.69259 | -81.896
0.6963 | -81.7401
0.7 | -81.7319
0.7037 | -81.9936
0.70741 | -82.0117
0.71111 | -81.625
0.71481 | -81.386
0.71852 | -81.688
0.72222 | -82.4361
0.72593 | -83.1565
0.72963 | -83.1781
0.73333 | -83.0129
0.73704 | -83.1269
0.74074 | -83.1437
0.74444 | -83.0547
0.74815 | -83.0288
0.75185 | -83.2233
0.75556 | -83.1639
0.75926 | -82.9857
0.76296 | -83.1792
0.76667 | -83.4917
0.77037 | -83.9774
0.77407 | -84.3005
0.77778 | -84.1284
0.78148 | -83.7775
0.78519 | -83.3347
0.78889 | -82.9075
0.79259 | -82.9993
0.7963 | -83.3591
0.8 | -83.6695
0.8037 | -83.7042
0.80741 | -83.5496
0.81111 | -83.4785
0.81481 | -83.8016
0.81852 | -84.209
0.82222 | -84.3079
0.82593 | -84.2424
0.82963 | -84.1576
0.83333 | -84.0977
0.83704 | -83.9351
0.84074 | -84.0463
0.84444 | -84.3995
0.84815 | -84.2082
0.85185 | -84.0994
0.85556 | -84.6069
0.85926 | -85.0822
0.86296 | -85.0632
0.86667 | -85.035
0.87037 | -84.6481
0.87407 | -84.0493
0.87778 | -84.0173
0.88148 | -84.211
0.88518 | -84.3253
0.88889 | -84.638
0.89259 | -85.0172
0.8963 | -85.2013
0.9 | -85.5435
0.9037 | -85.6793
0.90741 | -85.582
0.91111 | -85.2567
0.91482 | -84.6958
0.91852 | -84.3984
0.92222 | -84.4126
0.92593 | -84.6144
0.92963 | -84.8564
0.93333 | -85.2578
0.93704 | -85.5303
0.94074 | -85.7571
0.94444 | -85.9852
0.94815 | -85.9401
0.95185 | -85.7261
0.95556 | -85.5174
0.95926 | -85.3623
0.96296 | -85.3699
0.96667 | -85.2949
0.97037 | -85.1917
0.97407 | -85.4612
0.97778 | -85.9044
0.98148 | -86.2154
0.98518 | -86.4864
0.98889 | -86.6567
0.99259 | -86.6598
0.9963 | -86.5477
1 | -86.3778
1.0037 | -86.2084
1.0074 | -86.3037
1.0111 | -86.4503
1.0148 | -85.9735
1.0185 | -85.7243
1.0222 | -86.0827
1.0259 | -86.4509
1.0296 | -86.6376
1.0333 | -86.4558
1.037 | -85.9409
1.0407 | -85.7907
1.0444 | -85.984
1.0481 | -86.2019
1.0519 | -86.1538
1.0556 | -86.0571
1.0593 | -86.0611
1.063 | -85.9449
1.0667 | -85.916
1.0704 | -85.9831
1.0741 | -86.1533
1.0778 | -86.4151
1.0815 | -86.7237
1.0852 | -86.9283
1.0889 | -86.9862
1.0926 | -87.0517
1.0963 | -87.1728
1.1 | -87.236
1.1037 | -87.1514
1.1074 | -86.9387
1.1111 | -86.7765
1.1148 | -86.8008
1.1185 | -86.9214
1.1222 | -87.0154
1.1259 | -86.9982
1.1296 | -86.8569
1.1333 | -86.7875
1.137 | -86.9357
1.1407 | -87.14
1.1444 | -87.206
1.1481 | -87.2574
1.1519 | -87.4519
1.1556 | -87.682
1.1593 | -87.833
1.163 | -87.9017
1.1667 | -87.6681
1.1704 | -87.0323
1.1741 | -86.4042
1.1778 | -86.1021
1.1815 | -86.128
1.1852 | -86.3647
1.1889 | -86.6811
1.1926 | -87.0233
1.1963 | -87.3915
1.2 | -87.7728
1.2037 | -88.0451
1.2074 | -88.028
1.2111 | -87.7064
1.2148 | -87.4075
1.2185 | -87.417
1.2222 | -87.5764
1.2259 | -87.6747
1.2296 | -87.7845
1.2333 | -87.8656
1.237 | -87.7579
1.2407 | -87.4888
1.2444 | -87.3295
1.2481 | -87.4841
1.2519 | -87.8523
1.2556 | -88.152
1.2593 | -88.1933
1.263 | -88.0864
1.2667 | -88.0142
1.2704 | -87.974
1.2741 | -87.8435
1.2778 | -87.6456
1.2815 | -87.5876
1.2852 | -87.6816
1.2889 | -87.7584
1.2926 | -87.8581
1.2963 | -88.1065
1.3 | -88.4513
1.3037 | -88.6591
1.3074 | -88.6879
1.3111 | -88.7454
1.3148 | -88.8351
1.3185 | -88.9309
1.3222 | -89.1036
1.3259 | -89.12
1.3296 | -88.7966
1.3333 | -88.4002
1.337 | -88.1527
1.3407 | -88.0902
1.3444 | -88.1468
1.3481 | -88.226
1.3519 | -88.3439
1.3556 | -88.5608
1.3593 | -88.7144
1.363 | -88.495
1.3667 | -87.9961
1.3704 | -87.657
1.3741 | -87.7098
1.3778 | -88.089
1.3815 | -88.566
1.3852 | -88.9308
1.3889 | -89.0155
1.3926 | -88.8446
1.3963 | -88.7178
1.4 | -88.8325
1.4037 | -89.1258
1.4074 | -89.3878
1.4111 | -89.4725
1.4148 | -89.4096
1.4185 | -89.3421
1.4222 | -89.338
1.4259 | -89.2372
1.4296 | -88.996
1.4333 | -88.8229
1.437 | -88.7974
1.4407 | -88.9221
1.4444 | -89.2756
1.4481 | -89.822
1.4519 | -90.3501
1.4556 | -90.6586
1.4593 | -90.613
1.463 | -90.3246
1.4667 | -90.1608
1.4704 | -90.2772
1.4741 | -90.3539
1.4778 | -89.9183
1.4815 | -89.1481
1.4852 | -88.6269
1.4889 | -88.5818
1.4926 | -88.8713
1.4963 | -89.2328
1.5 | -89.5288
1.5037 | -89.7786
1.5074 | -90.0031
1.5111 | -90.1838
1.5148 | -90.2049
1.5185 | -89.9562
1.5222 | -89.5746
1.5259 | -89.3292
1.5296 | -89.4136
1.5333 | -89.7277
1.537 | -89.9049
1.5407 | -89.8411
1.5444 | -89.7992
1.5481 | -89.8851
1.5519 | -89.976
1.5556 | -90.078
1.5593 | -90.334
1.563 | -90.6761
1.5667 | -90.8774
1.5704 | -90.9033
1.5741 | -90.8572
1.5778 | -90.7062
1.5815 | -90.4439
1.5852 | -90.1295
1.5889 | -89.8371
1.5926 | -89.6973
1.5963 | -89.7911
1.6 | -90.1369
1.6037 | -90.4069
1.6074 | -90.5766
1.6111 | -90.5539
1.6148 | -90.3966
1.6185 | -90.2034
1.6222 | -90.0222
1.6259 | -89.9025
1.6296 | -89.914
1.6333 | -90.0654
1.637 | -90.2576
1.6407 | -90.3761
1.6444 | -90.3929
1.6481 | -90.3433
1.6519 | -90.2609
1.6556 | -90.2221
1.6593 | -90.2691
1.663 | -90.349
1.6667 | -90.4257
1.6704 | -90.5094
1.6741 | -90.6062
1.6778 | -90.6471
1.6815 | -90.5631
1.6852 | -90.3973
1.6889 | -90.2654
1.6926 | -90.2482
1.6963 | -90.3263
1.7 | -90.3843
1.7037 | -90.3892
1.7074 | -90.3823
1.7111 | -90.4609
1.7148 | -90.7018
1.7185 | -91.0849
1.7222 | -91.4553
1.7259 | -91.603
1.7296 | -91.5176
1.7333 | -91.3774
1.737 | -91.3067
1.7407 | -91.2991
1.7444 | -91.2817
1.7481 | -91.2219
1.7519 | -91.1302
1.7556 | -91.0353
1.7593 | -91.0282
1.763 | -91.1522
1.7667 | -91.3103
1.7704 | -91.3453
1.7741 | -91.2581
1.7778 | -91.1943
1.7815 | -91.2514
1.7852 | -91.338
1.7889 | -91.3163
1.7926 | -91.1854
1.7963 | -91.0644
1.8 | -91.0623
1.8037 | -91.1681
1.8074 | -91.2892
1.8111 | -91.3204
1.8148 | -91.2418
1.8185 | -91.1735
1.8222 | -91.2099
1.8259 | -91.3435
1.8296 | -91.4829
1.8333 | -91.4866
1.837 | -91.3431
1.8407 | -91.1544
1.8444 | -91.0092
1.8481 | -90.9335
1.8519 | -90.9261
1.8556 | -90.9924
1.8593 | -91.1178
1.863 | -91.2369
1.8667 | -91.2956
1.8704 | -91.2857
1.8741 | -91.2992
1.8778 | -91.4331
1.8815 | -91.7052
1.8852 | -92.0017
1.8889 | -92.1371
1.8926 | -91.9982
1.8963 | -91.7127
1.9 | -91.5228
1.9037 | -91.5288
1.9074 | -91.6388
1.9111 | -91.6902
1.9148 | -91.6006
1.9185 | -91.4391
1.9222 | -91.3164
1.9259 | -91.3031
1.9296 | -91.3873
1.9333 | -91.5473
1.937 | -91.7641
1.9407 | -91.9205
1.9444 | -91.9252
1.9481 | -91.8519
1.9519 | -91.8489
1.9556 | -91.9661
1.9593 | -92.1514
1.963 | -92.34
1.9667 | -92.4839
1.9704 | -92.5153
1.9741 | -92.392
1.9778 | -92.1414
1.9815 | -91.8748
1.9852 | -91.6841
1.9889 | -91.5906
1.9926 | -91.579
1.9963 | -91.6186
2 | -91.6249
2.0037 | -91.5248
2.0074 | -91.3686
2.0111 | -91.2919
2.0148 | -91.3506
2.0185 | -91.4874
2.0222 | -91.5984
2.0259 | -91.66
2.0296 | -91.731
2.0333 | -91.852
2.037 | -91.9947
2.0407 | -92.1146
2.0444 | -92.1837
2.0481 | -92.1576
2.0519 | -92.0781
2.0556 | -92.0797
2.0593 | -92.1977
2.063 | -92.3298
2.0667 | -92.3466
2.0704 | -92.2646
2.0741 | -92.2141
2.0778 | -92.2562
2.0815 | -92.3761
2.0852 | -92.5147
2.0889 | -92.6379
2.0926 | -92.687
2.0963 | -92.6441
2.1 | -92.5847
2.1037 | -92.5867
2.1074 | -92.6285
2.1111 | -92.6586
2.1148 | -92.5968
2.1185 | -92.4316
2.1222 | -92.2705
2.1259 | -92.2479
2.1296 | -92.3367
2.1333 | -92.3719
2.137 | -92.1598
2.1407 | -91.9997
2.1444 | -91.8604
2.1481 | -91.7797
2.1519 | -91.7753
2.1556 | -91.8341
2.1593 | -91.9391
2.163 | -92.0753
2.1667 | -92.1998
2.1704 | -92.2786
2.1741 | -92.3134
2.1778 | -92.3219
2.1815 | -92.3288
2.1852 | -92.3467
2.1889 | -92.3466
2.1926 | -92.2905
2.1963 | -92.1657
2.2 | -92.0205
2.2037 | -91.9281
2.2074 | -91.9536
2.2111 | -92.1126
2.2148 | -92.3829
2.2185 | -92.7136
2.2222 | -93.0084
2.2259 | -93.172
2.2296 | -93.1845
2.2333 | -93.1116
2.237 | -93.0622
2.2407 | -93.0786
2.2444 | -93.1586
2.2481 | -93.2623
2.2519 | -93.3321
2.2556 | -93.3541
2.2593 | -93.3161
2.263 | -93.2409
2.2667 | -93.1467
2.2704 | -93.0304
2.2741 | -92.8757
2.2778 | -92.6816
2.2815 | -92.5184
2.2852 | -92.4288
2.2889 | -92.3798
2.2926 | -92.323
2.2963 | -92.2055
2.3 | -92.0372
2.3037 | -91.8566
2.3074 | -91.7397
2.3111 | -91.7673
2.3148 | -91.9806
2.3185 | -92.3483
2.3222 | -92.7489
2.3259 | -93.062
2.3296 | -93.2184
2.3333 | -93.2178
2.337 | -93.1604
2.3407 | -93.1224
2.3444 | -93.1182
2.3481 | -93.1219
2.3519 | -93.1074
2.3556 | -93.0889
2.3593 | -93.0977
2.363 | -93.1536
2.3667 | -93.2252
2.3704 | -93.2771
2.3741 | -93.2719
2.3778 | -93.2077
2.3815 | -93.13
2.3852 | -93.0639
2.3889 | -93.031
2.3926 | -93.048
2.3963 | -93.1184
2.4 | -93.2244
2.4037 | -93.3109
2.4074 | -93.3427
2.4111 | -93.2923
2.4148 | -93.1647
2.4185 | -92.9741
2.4222 | -92.7753
2.4259 | -92.6607
2.4296 | -92.6761
2.4333 | -92.8158
2.437 | -93.0087
2.4407 | -93.157
2.4444 | -93.2165
2.4481 | -93.2051
2.4519 | -93.1914
2.4556 | -93.2098
2.4593 | -93.2828
2.463 | -93.3837
2.4667 | -93.4603
2.4704 | -93.4882
2.4741 | -93.4927
2.4778 | -93.4955
2.4815 | -93.516
2.4852 | -93.5374
2.4889 | -93.5348
2.4926 | -93.513
2.4963 | -93.5135
2.5 | -93.5303
2.5037 | -93.5513
2.5074 | -93.5474
2.5111 | -93.495
2.5148 | -93.4131
2.5185 | -93.3509
2.5222 | -93.3161
2.5259 | -93.3036
2.5296 | -93.313
2.5333 | -93.3494
2.537 | -93.4273
2.5407 | -93.556
2.5444 | -93.7226
2.5481 | -93.8902
2.5519 | -94.0062
2.5556 | -94.0453
2.5593 | -94.0304
2.563 | -93.9908
2.5667 | -93.9201
2.5704 | -93.8278
2.5741 | -93.7338
2.5778 | -93.6599
2.5815 | -93.6062
2.5852 | -93.5455
2.5889 | -93.4895
2.5926 | -93.4729
2.5963 | -93.5273
2.6 | -93.6572
2.6037 | -93.8328
2.6074 | -94.0041
2.6111 | -94.0923
2.6148 | -94.0809
2.6185 | -93.968
2.6222 | -93.8004
2.6259 | -93.6737
2.6296 | -93.6569
2.6333 | -93.7747
2.637 | -93.9781
2.6407 | -94.1887
2.6444 | -94.3405
2.6481 | -94.4044
2.6519 | -94.3856
2.6556 | -94.2848
2.6593 | -94.1515
2.663 | -94.0212
2.6667 | -93.9263
2.6704 | -93.9585
2.6741 | -93.9794
2.6778 | -94.0163
2.6815 | -94.0491
2.6852 | -94.0769
2.6889 | -94.1173
2.6926 | -94.2013
2.6963 | -94.3451
2.7 | -94.5389
2.7037 | -94.7408
2.7074 | -94.9143
2.7111 | -95.0028
2.7148 | -94.9847
2.7185 | -94.8707
2.7222 | -94.7215
2.7259 | -94.5949
2.7296 | -94.5243
2.7333 | -94.5212
2.737 | -94.5536
2.7407 | -94.5977
2.7444 | -94.6369
2.7481 | -94.6633
2.7519 | -94.6686
2.7556 | -94.646
2.7593 | -94.5873
2.763 | -94.4635
2.7667 | -94.3029
2.7704 | -94.1574
2.7741 | -94.0698
2.7778 | -94.0644
2.7815 | -94.136
2.7852 | -94.2696
2.7889 | -94.4224
2.7926 | -94.5661
2.7963 | -94.6782
2.8 | -94.7572
2.8037 | -94.7996
2.8074 | -94.8371
2.8111 | -94.9008
2.8148 | -94.9802
2.8185 | -95.0719
2.8222 | -95.1599
2.8259 | -95.2365
2.8296 | -95.3157
2.8333 | -95.3946
2.837 | -95.4374
2.8407 | -95.4169
2.8444 | -95.3228
2.8481 | -95.16
2.8519 | -94.971
2.8556 | -94.7904
2.8593 | -94.6494
2.863 | -94.5785
2.8667 | -94.5826
2.8704 | -94.65
2.8741 | -94.7386
2.8778 | -94.8058
2.8815 | -94.816
2.8852 | -94.7631
2.8889 | -94.6638
2.8926 | -94.569
2.8963 | -94.5238
2.9 | -94.5583
2.9037 | -94.6687
2.9074 | -94.8146
2.9111 | -94.9448
2.9148 | -94.9956
2.9185 | -94.948
2.9222 | -94.8198
2.9259 | -94.6466
2.9296 | -94.4852
2.9333 | -94.3708
2.937 | -94.319
2.9407 | -94.3383
2.9444 | -94.4285
2.9481 | -94.5755
2.9519 | -94.7686
2.9556 | -94.9926
2.9593 | -95.2101
2.963 | -95.3843
2.9667 | -95.4576
2.9704 | -95.4138
2.9741 | -95.2914
2.9778 | -95.1552
2.9815 | -95.0745
2.9852 | -95.0749
2.9889 | -95.1614
2.9926 | -95.2891
2.9963 | -95.4087
3 | -95.4914
3.0037 | -95.4962
3.0074 | -95.4343
3.0111 | -95.3446
3.0148 | -95.2702
3.0185 | -95.241
3.0222 | -95.2798
3.0259 | -95.3729
3.0296 | -95.4887
3.0333 | -95.5939
3.037 | -95.6491
3.0407 | -95.6578
3.0444 | -95.6285
3.0481 | -95.5839
3.0519 | -95.5579
3.0556 | -95.5366
3.0593 | -95.5163
3.063 | -95.49
3.0667 | -95.4565
3.0704 | -95.4302
3.0741 | -95.4321
3.0778 | -95.4666
3.0815 | -95.5246
3.0852 | -95.5962
3.0889 | -95.6623
3.0926 | -95.7128
3.0963 | -95.7387
3.1 | -95.7547
3.1037 | -95.7676
3.1074 | -95.7784
3.1111 | -95.7995
3.1148 | -95.8542
3.1185 | -95.9649
3.1222 | -96.1327
3.1259 | -96.3251
3.1296 | -96.4937
3.1333 | -96.5713
3.137 | -96.5293
3.1407 | -96.3755
3.1444 | -96.1618
3.1481 | -95.911
3.1519 | -95.6724
3.1556 | -95.4649
3.1593 | -95.3111
3.163 | -95.2238
3.1667 | -95.1967
3.1704 | -95.2247
3.1741 | -95.2981
3.1778 | -95.4144
3.1815 | -95.5868
3.1852 | -95.8085
3.1889 | -96.0376
3.1926 | -96.214
3.1963 | -96.3005
3.2 | -96.2586
3.2037 | -96.0233
3.2074 | -95.8597
3.2111 | -95.6833
3.2148 | -95.5111
3.2185 | -95.373
3.2222 | -95.2773
3.2259 | -95.2185
3.2296 | -95.1841
3.2333 | -95.1701
3.237 | -95.194
3.2407 | -95.2635
3.2444 | -95.379
3.2481 | -95.5169
3.2519 | -95.6477
3.2556 | -95.7484
3.2593 | -95.8041
3.263 | -95.8179
3.2667 | -95.8083
3.2704 | -95.8011
3.2741 | -95.7997
3.2778 | -95.8104
3.2815 | -95.8441
3.2852 | -95.9121
3.2889 | -96.0095
3.2926 | -96.1087
3.2963 | -96.1897
3.3 | -96.2284
3.3037 | -96.2256
3.3074 | -96.1944
3.3111 | -96.1517
3.3148 | -96.1182
3.3185 | -96.1101
3.3222 | -96.1385
3.3259 | -96.2151
3.3296 | -96.3109
3.3333 | -96.3989
3.337 | -96.4752
3.3407 | -96.5135
3.3444 | -96.5056
3.3481 | -96.4732
3.3519 | -96.4244
3.3556 | -96.3795
3.3593 | -96.339
3.363 | -96.3174
3.3667 | -96.3216
3.3704 | -96.354
3.3741 | -96.4136
3.3778 | -96.4852
3.3815 | -96.5554
3.3852 | -96.6372
3.3889 | -96.713
3.3926 | -96.7691
3.3963 | -96.822
3.4 | -96.8748
3.4037 | -96.9165
3.4074 | -96.9606
3.4111 | -96.9974
3.4148 | -97.0375
3.4185 | -97.0761
3.4222 | -97.1027
3.4259 | -97.1107
3.4296 | -97.0976
3.4333 | -97.0692
3.437 | -97.031
3.4407 | -96.9874
3.4444 | -96.9453
3.4481 | -96.8952
3.4519 | -96.8499
3.4556 | -96.8143
3.4593 | -96.7821
3.463 | -96.7599
3.4667 | -96.7425
3.4704 | -96.7061
3.4741 | -96.6394
3.4778 | -96.5398
3.4815 | -96.4062
3.4852 | -96.2472
3.4889 | -96.0814
3.4926 | -95.9257
3.4963 | -95.8069
3.5 | -95.7294
3.5037 | -95.6904
3.5074 | -95.6788
3.5111 | -95.6951
3.5148 | -95.7289
3.5185 | -95.7761
3.5222 | -95.8467
3.5259 | -95.9302
3.5296 | -96.0158
3.5333 | -96.0999
3.537 | -96.1845
3.5407 | -96.2825
3.5444 | -96.3958
3.5481 | -96.5287
3.5519 | -96.6644
3.5556 | -96.8038
3.5593 | -96.9215
3.563 | -96.9988
3.5667 | -97.0355
3.5704 | -97.0436
3.5741 | -97.0124
3.5778 | -96.9618
3.5815 | -96.9141
3.5852 | -96.8721
3.5889 | -96.8476
3.5926 | -96.8384
3.5963 | -96.8295
3.6 | -96.8004
3.6037 | -96.7458
3.6074 | -96.665
3.6111 | -96.5592
3.6148 | -96.4639
3.6185 | -96.4093
3.6222 | -96.4152
3.6259 | -96.4988
3.6296 | -96.6584
3.6333 | -96.8612
3.637 | -97.0653
3.6407 | -97.2339
3.6444 | -97.3398
3.6481 | -97.3906
3.6519 | -97.4171
3.6556 | -97.4403
3.6593 | -97.4643
3.663 | -97.4959
3.6667 | -97.5327
3.6704 | -97.5606
3.6741 | -97.5549
3.6778 | -97.513
3.6815 | -97.4512
3.6852 | -97.3991
3.6889 | -97.3832
3.6926 | -97.4176
3.6963 | -97.4812
3.7 | -97.5489
3.7037 | -97.5825
3.7074 | -97.5494
3.7111 | -97.448
3.7148 | -97.2997
3.7185 | -97.1374
3.7222 | -96.9902
3.7259 | -96.9001
3.7296 | -96.8978
3.7333 | -96.9686
3.737 | -97.1177
3.7407 | -97.2205
3.7444 | -97.3139
3.7481 | -97.3888
3.7519 | -97.4236
3.7556 | -97.4124
3.7593 | -97.3524
3.763 | -97.2546
3.7667 | -97.1382
3.7704 | -97.0292
3.7741 | -96.943
3.7778 | -96.8992
3.7815 | -96.9143
3.7852 | -96.9735
3.7889 | -97.0557
3.7926 | -97.149
3.7963 | -97.2388
3.8 | -97.3163
3.8037 | -97.3815
3.8074 | -97.4298
3.8111 | -97.4427
3.8148 | -97.4488
3.8185 | -97.4544
3.8222 | -97.4492
3.8259 | -97.4319
3.8296 | -97.4143
3.8333 | -97.3992
3.837 | -97.3977
3.8407 | -97.4224
3.8444 | -97.4674
3.8481 | -97.5332
3.8519 | -97.6158
3.8556 | -97.6841
3.8593 | -97.7206
3.863 | -97.7256
3.8667 | -97.7006
3.8704 | -97.6695
3.8741 | -97.6466
3.8778 | -97.6275
3.8815 | -97.6125
3.8852 | -97.6211
3.8889 | -97.6391
3.8926 | -97.6593
3.8963 | -97.6801
3.9 | -97.6996
3.9037 | -97.7082
3.9074 | -97.7127
3.9111 | -97.7124
3.9148 | -97.7178
3.9185 | -97.7323
3.9222 | -97.7339
3.9259 | -97.7208
3.9296 | -97.6955
3.9333 | -97.6618
3.937 | -97.6258
3.9407 | -97.5951
3.9444 | -97.5782
3.9481 | -97.5716
3.9519 | -97.5763
3.9556 | -97.5799
3.9593 | -97.5738
3.963 | -97.5494
3.9667 | -97.5107
3.9704 | -97.4618
3.9741 | -97.4019
3.9778 | -97.3379
3.9815 | -97.2843
3.9852 | -97.2442
3.9889 | -97.2282
3.9926 | -97.2501
3.9963 | -97.3134
4 | -97.4245
4.0037 | -97.5673
4.0074 | -97.7292
4.0111 | -97.897
4.0148 | -98.0511
4.0185 | -98.1737
4.0222 | -98.2587
4.0259 | -98.3055
4.0296 | -98.3078
4.0333 | -98.2799
4.037 | -98.2427
4.0407 | -98.2048
4.0444 | -98.1729
4.0481 | -98.1671
4.0519 | -98.1833
4.0556 | -98.2018
4.0593 | -98.2272
4.063 | -98.2553
4.0667 | -98.2878
4.0704 | -98.3098
4.0741 | -98.3205
4.0778 | -98.3361
4.0815 | -98.3519
4.0852 | -98.3759
4.0889 | -98.4102
4.0926 | -98.4549
4.0963 | -98.4881
4.1 | -98.5009
4.1037 | -98.4939
4.1074 | -98.4678
4.1111 | -98.4194
4.1148 | -98.3619
4.1185 | -98.3218
4.1222 | -98.3005
4.1259 | -98.2984
4.1296 | -98.3115
4.1333 | -98.3403
4.137 | -98.3627
4.1407 | -98.3649
4.1444 | -98.3485
4.1481 | -98.2992
4.1519 | -98.2284
4.1556 | -98.1472
4.1593 | -98.0701
4.163 | -98.0055
4.1667 | -97.9568
4.1704 | -97.9424
4.1741 | -97.9533
4.1778 | -97.9845
4.1815 | -98.0175
4.1852 | -98.0463
4.1889 | -98.058
4.1926 | -98.0393
4.1963 | -97.9853
4.2 | -97.9101
4.2037 | -97.8309
4.2074 | -97.7552
4.2111 | -97.697
4.2148 | -97.6563
4.2185 | -97.6444
4.2222 | -97.6655
4.2259 | -97.7226
4.2296 | -97.8167
4.2333 | -97.934
4.237 | -98.0777
4.2407 | -98.2284
4.2444 | -98.3706
4.2481 | -98.4881
4.2519 | -98.5788
4.2556 | -98.6383
4.2593 | -98.6721
4.263 | -98.679
4.2667 | -98.6639
4.2704 | -98.6348
4.2741 | -98.58
4.2778 | -98.5331
4.2815 | -98.4764
4.2852 | -98.4077
4.2889 | -98.3306
4.2926 | -98.2573
4.2963 | -98.1865
4.3 | -98.1282
4.3037 | -98.0805
4.3074 | -98.0387
4.3111 | -98.0082
4.3148 | -97.9922
4.3185 | -97.9887
4.3222 | -97.9949
4.3259 | -98.0051
4.3296 | -98.0165
4.3333 | -98.0277
4.337 | -98.0427
4.3407 | -98.0523
4.3444 | -98.0485
4.3481 | -98.0418
4.3519 | -98.0346
4.3556 | -98.0306
4.3593 | -98.0333
4.363 | -98.0366
4.3667 | -98.0576
4.3704 | -98.0938
4.3741 | -98.1421
4.3778 | -98.2068
4.3815 | -98.278
4.3852 | -98.3473
4.3889 | -98.4245
4.3926 | -98.5037
4.3963 | -98.5863
4.4 | -98.6751
4.4037 | -98.7593
4.4074 | -98.8319
4.4111 | -98.8983
4.4148 | -98.9508
4.4185 | -98.9861
4.4222 | -98.9992
4.4259 | -98.994
4.4296 | -98.9804
4.4333 | -98.9575
4.437 | -98.9241
4.4407 | -98.8836
4.4444 | -98.8382
4.4481 | -98.7926
4.4519 | -98.7519
4.4556 | -98.7152
4.4593 | -98.6824
4.463 | -98.6598
4.4667 | -98.6356
4.4704 | -98.6009
4.4741 | -98.5755
4.4778 | -98.5534
4.4815 | -98.5375
4.4852 | -98.5302
4.4889 | -98.5362
4.4926 | -98.5565
4.4963 | -98.5877
4.5 | -98.6347
4.5037 | -98.7055
4.5074 | -98.7934
4.5111 | -98.8966
4.5148 | -99.0093
4.5185 | -99.1136
4.5222 | -99.2036
4.5259 | -99.2761
4.5296 | -99.3233
4.5333 | -99.3452
4.537 | -99.354
4.5407 | -99.3595
4.5444 | -99.3453
4.5481 | -99.3283
4.5519 | -99.3049
4.5556 | -99.276
4.5593 | -99.2388
4.563 | -99.1914
4.5667 | -99.1449
4.5704 | -99.0925
4.5741 | -99.0408
4.5778 | -98.994
4.5815 | -98.9491
4.5852 | -98.9057
4.5889 | -98.8704
4.5926 | -98.8468
4.5963 | -98.8336
4.6 | -98.8347
4.6037 | -98.8461
4.6074 | -98.8584
4.6111 | -98.8686
4.6148 | -98.8764
4.6185 | -98.8748
4.6222 | -98.8697
4.6259 | -98.8574
4.6296 | -98.8393
4.6333 | -98.8219
4.637 | -98.8028
4.6407 | -98.7895
4.6444 | -98.7882
4.6481 | -98.8051
4.6519 | -98.8437
4.6556 | -98.8969
4.6593 | -98.9574
4.663 | -99.0237
4.6667 | -99.0899
4.6704 | -99.154
4.6741 | -99.214
4.6778 | -99.2637
4.6815 | -99.3025
4.6852 | -99.325
4.6889 | -99.3309
4.6926 | -99.3075
4.6963 | -99.2733
4.7 | -99.231
4.7037 | -99.1829
4.7074 | -99.1316
4.7111 | -99.0868
4.7148 | -99.0538
4.7185 | -99.0362
4.7222 | -99.0422
4.7259 | -99.0709
4.7296 | -99.1132
4.7333 | -99.1589
4.737 | -99.1957
4.7407 | -99.2021
4.7444 | -99.1839
4.7481 | -99.1515
4.7519 | -99.1025
4.7556 | -99.0475
4.7593 | -98.9999
4.763 | -98.961
4.7667 | -98.9218
4.7704 | -98.8966
4.7741 | -98.8859
4.7778 | -98.8832
4.7815 | -98.8838
4.7852 | -98.8881
4.7889 | -98.9007
4.7926 | -98.9197
4.7963 | -98.9387
4.8 | -98.9657
4.8037 | -98.9948
4.8074 | -99.0245
4.8111 | -99.0509
4.8148 | -99.0752
4.8185 | -99.0869
4.8222 | -99.0874
4.8259 | -99.0839
4.8296 | -99.0764
4.8333 | -99.0638
4.837 | -99.0537
4.8407 | -99.0437
4.8444 | -99.0295
4.8481 | -99.0291
4.8519 | -99.0491
4.8556 | -99.0882
4.8593 | -99.1425
4.863 | -99.2071
4.8667 | -99.2844
4.8704 | -99.3765
4.8741 | -99.4664
4.8778 | -99.5585
4.8815 | -99.6506
4.8852 | -99.7415
4.8889 | -99.8251
4.8926 | -99.9015
4.8963 | -99.9643
4.9 | -100.0058
4.9037 | -100.0251
4.9074 | -100.0252
4.9111 | -100.0042
4.9148 | -99.967
4.9185 | -99.9246
4.9222 | -99.8744
4.9259 | -99.8061
4.9296 | -99.726
4.9333 | -99.6352
4.937 | -99.5335
4.9407 | -99.4329
4.9444 | -99.3312
4.9481 | -99.2397
4.9519 | -99.1616
4.9556 | -99.1025
4.9593 | -99.0659
4.963 | -99.0528
4.9667 | -99.0687
4.9704 | -99.1061
4.9741 | -99.1598
4.9778 | -99.2226
4.9815 | -99.2867
4.9852 | -99.346
4.9889 | -99.3962
4.9926 | -99.4364
4.9963 | -99.4565
5 | -99.4685
5.0037 | -99.4815
5.0074 | -99.4942
5.0111 | -99.5023
5.0148 | -99.5104
5.0185 | -99.5161
5.0222 | -99.5263
5.0259 | -99.535
5.0296 | -99.5557
5.0333 | -99.5834
5.037 | -99.6168
5.0407 | -99.6559
5.0444 | -99.6942
5.0481 | -99.7316
5.0519 | -99.7656
5.0556 | -99.7992
5.0593 | -99.8248
5.063 | -99.8405
5.0667 | -99.8435
5.0704 | -99.8377
5.0741 | -99.8289
5.0778 | -99.8181
5.0815 | -99.8042
5.0852 | -99.7945
5.0889 | -99.8011
5.0926 | -99.8185
5.0963 | -99.8415
5.1 | -99.8774
5.1037 | -99.924
5.1074 | -99.9697
5.1111 | -100.008
5.1148 | -100.0447
5.1185 | -100.078
5.1222 | -100.1097
5.1259 | -100.1386
5.1296 | -100.1565
5.1333 | -100.1582
5.137 | -100.1461
5.1407 | -100.1201
5.1444 | -100.0866
5.1481 | -100.048
5.1519 | -100.011
5.1556 | -99.9836
5.1593 | -99.9778
5.163 | -99.9886
5.1667 | -100.0154
5.1704 | -100.0577
5.1741 | -100.1118
5.1778 | -100.1679
5.1815 | -100.2233
5.1852 | -100.2752
5.1889 | -100.3153
5.1926 | -100.3294
5.1963 | -100.3165
5.2 | -100.2765
5.2037 | -100.2147
5.2074 | -100.1445
5.2111 | -100.073
5.2148 | -100.0039
5.2185 | -99.9452
5.2222 | -99.901
5.2259 | -99.868
5.2296 | -99.8515
5.2333 | -99.8479
5.237 | -99.8574
5.2407 | -99.8871
5.2444 | -99.9307
5.2481 | -99.9883
5.2519 | -100.065
5.2556 | -100.1533
5.2593 | -100.2491
5.263 | -100.3416
5.2667 | -100.4331
5.2704 | -100.5171
5.2741 | -100.5842
5.2778 | -100.6322
5.2815 | -100.6562
5.2852 | -100.6479
5.2889 | -100.6133
5.2926 | -100.5478
5.2963 | -100.4563
5.3 | -100.3514
5.3037 | -100.2422
5.3074 | -100.1376
5.3111 | -100.0497
5.3148 | -99.9836
5.3185 | -99.9414
5.3222 | -99.9165
5.3259 | -99.9143
5.3296 | -99.929
5.3333 | -99.9595
5.337 | -100.0043
5.3407 | -100.0536
5.3444 | -100.1042
5.3481 | -100.1543
5.3519 | -100.2029
5.3556 | -100.2432
5.3593 | -100.2775
5.363 | -100.3071
5.3667 | -100.3281
5.3704 | -100.3352
5.3741 | -100.3361
5.3778 | -100.324
5.3815 | -100.3056
5.3852 | -100.2884
5.3889 | -100.2719
5.3926 | -100.2576
5.3963 | -100.2459
5.4 | -100.2426
5.4037 | -100.2627
5.4074 | -100.2937
5.4111 | -100.3389
5.4148 | -100.3945
5.4185 | -100.4565
5.4222 | -100.5211
5.4259 | -100.5712
5.4296 | -100.6075
5.4333 | -100.632
5.437 | -100.6399
5.4407 | -100.6377
5.4444 | -100.6179
5.4481 | -100.5901
5.4519 | -100.5622
5.4556 | -100.5341
5.4593 | -100.506
5.463 | -100.4831
5.4667 | -100.4577
5.4704 | -100.4302
5.4741 | -100.4049
5.4778 | -100.3831
5.4815 | -100.3654
5.4852 | -100.3458
5.4889 | -100.3212
5.4926 | -100.3011
5.4963 | -100.2857
5.5 | -100.2811
5.5037 | -100.2804
5.5074 | -100.2905
5.5111 | -100.314
5.5148 | -100.3378
5.5185 | -100.3645
5.5222 | -100.3929
5.5259 | -100.4184
5.5296 | -100.4449
5.5333 | -100.4647
5.537 | -100.4753
5.5407 | -100.4783
5.5444 | -100.4794
5.5481 | -100.4769
5.5519 | -100.4729
5.5556 | -100.4671
5.5593 | -100.4553
5.563 | -100.4419
5.5667 | -100.4264
5.5704 | -100.4115
5.5741 | -100.4025
5.5778 | -100.3974
5.5815 | -100.3976
5.5852 | -100.4045
5.5889 | -100.4296
5.5926 | -100.4645
5.5963 | -100.4997
5.6 | -100.5449
5.6037 | -100.5877
5.6074 | -100.6234
5.6111 | -100.6546
5.6148 | -100.6722
5.6185 | -100.6714
5.6222 | -100.664
5.6259 | -100.6498
5.6296 | -100.6297
5.6333 | -100.6089
5.637 | -100.5973
5.6407 | -100.5956
5.6444 | -100.6038
5.6481 | -100.6256
5.6519 | -100.6545
5.6556 | -100.6859
5.6593 | -100.7176
5.663 | -100.7492
5.6667 | -100.7775
5.6704 | -100.7997
5.6741 | -100.8091
5.6778 | -100.8132
5.6815 | -100.8176
5.6852 | -100.8147
5.6889 | -100.8094
5.6926 | -100.7974
5.6963 | -100.7774
5.7 | -100.7501
5.7037 | -100.7179
5.7074 | -100.6732
5.7111 | -100.6227
5.7148 | -100.5735
5.7185 | -100.5326
5.7222 | -100.5023
5.7259 | -100.4867
5.7296 | -100.479
5.7333 | -100.472
5.737 | -100.4711
5.7407 | -100.4711
5.7444 | -100.4719
5.7481 | -100.4752
5.7519 | -100.4843
5.7556 | -100.4982
5.7593 | -100.5207
5.763 | -100.5448
5.7667 | -100.574
5.7704 | -100.6076
5.7741 | -100.6396
5.7778 | -100.6688
5.7815 | -100.6897
5.7852 | -100.7043
5.7889 | -100.7155
5.7926 | -100.7291
5.7963 | -100.7418
5.8 | -100.7539
5.8037 | -100.7659
5.8074 | -100.7764
5.8111 | -100.7904
5.8148 | -100.8126
5.8185 | -100.8356
5.8222 | -100.8585
5.8259 | -100.8743
5.8296 | -100.8776
5.8333 | -100.8715
5.837 | -100.8535
5.8407 | -100.8276
5.8444 | -100.7887
5.8481 | -100.7455
5.8519 | -100.6934
5.8556 | -100.6379
5.8593 | -100.5882
5.863 | -100.5428
5.8667 | -100.5076
5.8704 | -100.4854
5.8741 | -100.5005
5.8778 | -100.5053
5.8815 | -100.5151
5.8852 | -100.5326
5.8889 | -100.5615
5.8926 | -100.5969
5.8963 | -100.6352
5.9 | -100.6743
5.9037 | -100.7149
5.9074 | -100.7531
5.9111 | -100.7837
5.9148 | -100.8037
5.9185 | -100.8188
5.9222 | -100.8208
5.9259 | -100.8166
5.9296 | -100.8035
5.9333 | -100.7829
5.937 | -100.7571
5.9407 | -100.7338
5.9444 | -100.7104
5.9481 | -100.6898
5.9519 | -100.6784
5.9556 | -100.6701
5.9593 | -100.6627
5.963 | -100.6615
5.9667 | -100.6719
5.9704 | -100.6962
5.9741 | -100.7292
5.9778 | -100.7706
5.9815 | -100.8141
5.9852 | -100.8614
5.9889 | -100.9121
5.9926 | -100.9548
5.9963 | -100.9903
6 | -101.0227
6.0037 | -101.0523
6.0074 | -101.0795
6.0111 | -101.0964
6.0148 | -101.1123
6.0185 | -101.1211
6.0222 | -101.1324
6.0259 | -101.1436
6.0296 | -101.1516
6.0333 | -101.1496
6.037 | -101.1439
6.0407 | -101.1412
6.0444 | -101.1357
6.0481 | -101.125
6.0519 | -101.1098
6.0556 | -101.0906
6.0593 | -101.0623
6.063 | -101.0312
6.0667 | -101.0141
6.0704 | -100.9986
6.0741 | -100.9829
6.0778 | -100.9728
6.0815 | -100.9713
6.0852 | -100.9769
6.0889 | -100.9829
6.0926 | -100.9903
6.0963 | -100.9947
6.1 | -101.0008
6.1037 | -101.0044
6.1074 | -101.0023
6.1111 | -100.9986
6.1148 | -100.9916
6.1185 | -100.9881
6.1222 | -100.9857
6.1259 | -100.9868
6.1296 | -100.9894
6.1333 | -100.9921
6.137 | -101.0023
6.1407 | -101.0185
6.1444 | -101.0385
6.1481 | -101.0563
6.1519 | -101.0756
6.1556 | -101.0905
6.1593 | -101.0932
6.163 | -101.0922
6.1667 | -101.0838
6.1704 | -101.069
6.1741 | -101.0492
6.1778 | -101.0317
6.1815 | -101.0173
6.1852 | -101.0077
6.1889 | -101.0028
6.1926 | -101.0007
6.1963 | -101.0099
6.2 | -101.0323
6.2037 | -101.0577
6.2074 | -101.0884
6.2111 | -101.1231
6.2148 | -101.1628
6.2185 | -101.2019
6.2222 | -101.2341
6.2259 | -101.254
6.2296 | -101.2694
6.2333 | -101.2828
6.237 | -101.2895
6.2407 | -101.2921
6.2444 | -101.2926
6.2481 | -101.2907
6.2519 | -101.2899
6.2556 | -101.2936
6.2593 | -101.3026
6.263 | -101.3093
6.2667 | -101.3063
6.2704 | -101.3043
6.2741 | -101.3058
6.2778 | -101.2942
6.2815 | -101.2781
6.2852 | -101.2581
6.2889 | -101.2324
6.2926 | -101.2018
6.2963 | -101.1686
6.3 | -101.1283
6.3037 | -101.0928
6.3074 | -101.0627
6.3111 | -101.0453
6.3148 | -101.0431
6.3185 | -101.0543
6.3222 | -101.0779
6.3259 | -101.1123
6.3296 | -101.1522
6.3333 | -101.1921
6.337 | -101.2237
6.3407 | -101.248
6.3444 | -101.2652
6.3481 | -101.2669
6.3519 | -101.2496
6.3556 | -101.2219
6.3593 | -101.1977
6.363 | -101.1628
6.3667 | -101.1278
6.3704 | -101.1045
6.3741 | -101.0856
6.3778 | -101.0746
6.3815 | -101.0782
6.3852 | -101.1002
6.3889 | -101.1381
6.3926 | -101.1797
6.3963 | -101.2226
6.4 | -101.2698
6.4037 | -101.3087
6.4074 | -101.3351
6.4111 | -101.3409
6.4148 | -101.3588
6.4185 | -101.3691
6.4222 | -101.3786
6.4259 | -101.3934
6.4296 | -101.4088
6.4333 | -101.4292
6.437 | -101.4536
6.4407 | -101.4783
6.4444 | -101.5096
6.4481 | -101.5499
6.4519 | -101.5919
6.4556 | -101.6337
6.4593 | -101.6836
6.463 | -101.7328
6.4667 | -101.783
6.4704 | -101.8333
6.4741 | -101.8839
6.4778 | -101.9285
6.4815 | -101.9656
6.4852 | -102.0025
6.4889 | -102.0359
6.4926 | -102.0524
6.4963 | -102.0522
6.5 | -102.0363
6.5037 | -102.0008
6.5074 | -101.9457
6.5111 | -101.8793
6.5148 | -101.8021
6.5185 | -101.7192
6.5222 | -101.6364
6.5259 | -101.5577
6.5296 | -101.4897
6.5333 | -101.4334
6.537 | -101.383
6.5407 | -101.3408
6.5444 | -101.3099
6.5481 | -101.2882
6.5519 | -101.2753
6.5556 | -101.2701
6.5593 | -101.2662
6.563 | -101.2612
6.5667 | -101.2537
6.5704 | -101.2456
6.5741 | -101.2372
6.5778 | -101.2252
6.5815 | -101.2151
6.5852 | -101.206
6.5889 | -101.1924
6.5926 | -101.1838
6.5963 | -101.1805
6.6 | -101.1777
6.6037 | -101.1804
6.6074 | -101.1841
6.6111 | -101.1934
6.6148 | -101.2112
6.6185 | -101.2329
6.6222 | -101.2574
6.6259 | -101.2908
6.6296 | -101.3305
6.6333 | -101.3735
6.637 | -101.4206
6.6407 | -101.4646
6.6444 | -101.5026
6.6481 | -101.5323
6.6519 | -101.5566
6.6556 | -101.5741
6.6593 | -101.5944
6.663 | -101.6099
6.6667 | -101.6242
6.6704 | -101.6424
6.6741 | -101.6629
6.6778 | -101.6917
6.6815 | -101.7297
6.6852 | -101.7721
6.6889 | -101.8289
6.6926 | -101.8811
6.6963 | -101.9242
6.7 | -101.9709
6.7037 | -102.0063
6.7074 | -102.0231
6.7111 | -102.0169
6.7148 | -101.9951
6.7185 | -101.9648
6.7222 | -101.9264
6.7259 | -101.8766
6.7296 | -101.8233
6.7333 | -101.7782
6.737 | -101.7467
6.7407 | -101.7286
6.7444 | -101.7221
6.7481 | -101.7289
6.7519 | -101.7425
6.7556 | -101.7656
6.7593 | -101.7938
6.763 | -101.8313
6.7667 | -101.8675
6.7704 | -101.8928
6.7741 | -101.9105
6.7778 | -101.9197
6.7815 | -101.9194
6.7852 | -101.9156
6.7889 | -101.9025
6.7926 | -101.8909
6.7963 | -101.8848
6.8 | -101.8747
6.8037 | -101.8665
6.8074 | -101.8641
6.8111 | -101.8688
6.8148 | -101.8818
6.8185 | -101.8967
6.8222 | -101.9141
6.8259 | -101.9367
6.8296 | -101.9583
6.8333 | -101.9711
6.837 | -101.9745
6.8407 | -101.9744
6.8444 | -101.9722
6.8481 | -101.9694
6.8519 | -101.9663
6.8556 | -101.9617
6.8593 | -101.9625
6.863 | -101.9633
6.8667 | -101.9681
6.8704 | -101.9784
6.8741 | -101.9858
6.8778 | -101.9955
6.8815 | -101.9989
6.8852 | -101.999
6.8889 | -102.0004
6.8926 | -102.0013
6.8963 | -101.9962
6.9 | -101.9903
6.9037 | -101.9805
6.9074 | -101.9647
6.9111 | -101.9487
6.9148 | -101.9218
6.9185 | -101.8832
6.9222 | -101.8424
6.9259 | -101.8006
6.9296 | -101.7592
6.9333 | -101.7181
6.937 | -101.6795
6.9407 | -101.6434
6.9444 | -101.629
6.9481 | -101.6107
6.9519 | -101.5974
6.9556 | -101.5896
6.9593 | -101.5914
6.963 | -101.5991
6.9667 | -101.618
6.9704 | -101.6434
6.9741 | -101.675
6.9778 | -101.7118
6.9815 | -101.7457
6.9852 | -101.7815
6.9889 | -101.8172
6.9926 | -101.8528
6.9963 | -101.8809
7 | -101.9081
7.0037 | -101.934
7.0074 | -101.9553
7.0111 | -101.9752
7.0148 | -101.9809
7.0185 | -101.9831
7.0222 | -101.9881
7.0259 | -101.9876
7.0296 | -101.9814
7.0333 | -101.9724
7.037 | -101.9629
7.0407 | -101.9521
7.0444 | -101.9408
7.0481 | -101.929
7.0519 | -101.9262
7.0556 | -101.9262
7.0593 | -101.9344
7.063 | -101.9518
7.0667 | -101.9696
7.0704 | -101.9899
7.0741 | -102.0115
7.0778 | -102.0364
7.0815 | -102.0606
7.0852 | -102.0808
7.0889 | -102.1021
7.0926 | -102.122
7.0963 | -102.1384
7.1 | -102.1534
7.1037 | -102.1722
7.1074 | -102.1882
7.1111 | -102.1997
7.1148 | -102.2149
7.1185 | -102.2283
7.1222 | -102.2366
7.1259 | -102.2429
7.1296 | -102.247
7.1333 | -102.2441
7.137 | -102.2351
7.1407 | -102.224
7.1444 | -102.213
7.1481 | -102.196
7.1519 | -102.1765
7.1556 | -102.1514
7.1593 | -102.1166
7.163 | -102.0853
7.1667 | -102.0565
7.1704 | -102.0318
7.1741 | -102.0166
7.1778 | -102.0091
7.1815 | -102.0134
7.1852 | -102.0289
7.1889 | -102.0566
7.1926 | -102.0975
7.1963 | -102.1392
7.2 | -102.1834
7.2037 | -102.228
7.2074 | -102.2668
7.2111 | -102.2984
7.2148 | -102.3261
7.2185 | -102.3473
7.2222 | -102.3506
7.2259 | -102.3474
7.2296 | -102.3414
7.2333 | -102.3265
7.237 | -102.3033
7.2407 | -102.2771
7.2444 | -102.2546
7.2481 | -102.226
7.2519 | -102.1965
7.2556 | -102.1717
7.2593 | -102.1457
7.263 | -102.1271
7.2667 | -102.113
7.2704 | -102.1008
7.2741 | -102.0928
7.2778 | -102.0901
7.2815 | -102.0935
7.2852 | -102.1039
7.2889 | -102.1156
7.2926 | -102.1323
7.2963 | -102.1545
7.3 | -102.1808
7.3037 | -102.2156
7.3074 | -102.2532
7.3111 | -102.2939
7.3148 | -102.3421
7.3185 | -102.3948
7.3222 | -102.4501
7.3259 | -102.5063
7.3296 | -102.5682
7.3333 | -102.6272
7.337 | -102.6857
7.3407 | -102.7407
7.3444 | -102.7931
7.3481 | -102.8431
7.3519 | -102.8834
7.3556 | -102.918
7.3593 | -102.942
7.363 | -102.9508
7.3667 | -102.9494
7.3704 | -102.9328
7.3741 | -102.9094
7.3778 | -102.8805
7.3815 | -102.8505
7.3852 | -102.8162
7.3889 | -102.7834
7.3926 | -102.7578
7.3963 | -102.7425
7.4 | -102.7302
7.4037 | -102.7312
7.4074 | -102.7427
7.4111 | -102.7536
7.4148 | -102.7678
7.4185 | -102.784
7.4222 | -102.7964
7.4259 | -102.8028
7.4296 | -102.7982
7.4333 | -102.7843
7.437 | -102.7652
7.4407 | -102.738
7.4444 | -102.6973
7.4481 | -102.6506
7.4519 | -102.5984
7.4556 | -102.5373
7.4593 | -102.464
7.463 | -102.3908
7.4667 | -102.3154
7.4704 | -102.2465
7.4741 | -102.1821
7.4778 | -102.1412
7.4815 | -102.0933
7.4852 | -102.0516
7.4889 | -102.017
7.4926 | -101.9964
7.4963 | -101.9933
7.5 | -102.0011
7.5037 | -102.0182
7.5074 | -102.0451
7.5111 | -102.0801
7.5148 | -102.1188
7.5185 | -102.1625
7.5222 | -102.2105
7.5259 | -102.2587
7.5296 | -102.3032
7.5333 | -102.3427
7.537 | -102.3802
7.5407 | -102.4164
7.5444 | -102.4472
7.5481 | -102.4724
7.5519 | -102.4918
7.5556 | -102.5058
7.5593 | -102.5119
7.563 | -102.5151
7.5667 | -102.5162
7.5704 | -102.516
7.5741 | -102.5034
7.5778 | -102.489
7.5815 | -102.4753
7.5852 | -102.4592
7.5889 | -102.4435
7.5926 | -102.4322
7.5963 | -102.4269
7.6 | -102.4292
7.6037 | -102.438
7.6074 | -102.4486
7.6111 | -102.4628
7.6148 | -102.4814
7.6185 | -102.5009
7.6222 | -102.5235
7.6259 | -102.549
7.6296 | -102.5719
7.6333 | -102.5942
7.637 | -102.6153
7.6407 | -102.634
7.6444 | -102.6478
7.6481 | -102.6607
7.6519 | -102.6744
7.6556 | -102.6886
7.6593 | -102.7057
7.663 | -102.7188
7.6667 | -102.7285
7.6704 | -102.745
7.6741 | -102.7571
7.6778 | -102.7604
7.6815 | -102.7643
7.6852 | -102.7658
7.6889 | -102.766
7.6926 | -102.7559
7.6963 | -102.7397
7.7 | -102.7277
7.7037 | -102.7177
7.7074 | -102.707
7.7111 | -102.6983
7.7148 | -102.6948
7.7185 | -102.6989
7.7222 | -102.7061
7.7259 | -102.7206
7.7296 | -102.7469
7.7333 | -102.7809
7.737 | -102.8116
7.7407 | -102.8429
7.7444 | -102.8786
7.7481 | -102.9104
7.7519 | -102.939
7.7556 | -102.9619
7.7593 | -102.9799
7.763 | -102.9906
7.7667 | -102.9918
7.7704 | -102.983
7.7741 | -102.9655
7.7778 | -102.9485
7.7815 | -102.931
7.7852 | -102.9073
7.7889 | -102.8849
7.7926 | -102.8664
7.7963 | -102.8451
7.8 | -102.8294
7.8037 | -102.818
7.8074 | -102.8151
7.8111 | -102.8152
7.8148 | -102.8175
7.8185 | -102.8245
7.8222 | -102.8345
7.8259 | -102.8464
7.8296 | -102.8578
7.8333 | -102.8667
7.837 | -102.8773
7.8407 | -102.8846
7.8444 | -102.8905
7.8481 | -102.8924
7.8519 | -102.896
7.8556 | -102.8996
7.8593 | -102.8976
7.863 | -102.8942
7.8667 | -102.8824
7.8704 | -102.8647
7.8741 | -102.8466
7.8778 | -102.8248
7.8815 | -102.7967
7.8852 | -102.7708
7.8889 | -102.7483
7.8926 | -102.7203
7.8963 | -102.6919
7.9 | -102.6652
7.9037 | -102.6441
7.9074 | -102.6255
7.9111 | -102.6127
7.9148 | -102.6014
7.9185 | -102.5882
7.9222 | -102.577
7.9259 | -102.5676
7.9296 | -102.5647
7.9333 | -102.5652
7.937 | -102.5639
7.9407 | -102.5663
7.9444 | -102.572
7.9481 | -102.5811
7.9519 | -102.5849
7.9556 | -102.586
7.9593 | -102.5823
7.963 | -102.5779
7.9667 | -102.5706
7.9704 | -102.5572
7.9741 | -102.5445
7.9778 | -102.5326
7.9815 | -102.5259
7.9852 | -102.519
7.9889 | -102.5057
7.9926 | -102.4932
7.9963 | -102.4914
8 | -102.4903
8.0037 | -102.4896
8.0074 | -102.497
8.0111 | -102.5132
8.0148 | -102.5477
8.0185 | -102.5743
8.0222 | -102.6065
8.0259 | -102.6477
8.0296 | -102.6956
8.0333 | -102.749
8.037 | -102.8104
8.0407 | -102.8712
8.0444 | -102.9337
8.0481 | -102.9929
8.0519 | -103.0551
8.0556 | -103.1219
8.0593 | -103.1852
8.063 | -103.2428
8.0667 | -103.2911
8.0704 | -103.3345
8.0741 | -103.368
8.0778 | -103.3969
8.0815 | -103.4217
8.0852 | -103.4357
8.0889 | -103.4388
8.0926 | -103.4367
8.0963 | -103.4329
8.1 | -103.4275
8.1037 | -103.4223
8.1074 | -103.4213
8.1111 | -103.4135
8.1148 | -103.4041
8.1185 | -103.3998
8.1222 | -103.4003
8.1259 | -103.402
8.1296 | -103.4022
8.1333 | -103.4047
8.137 | -103.4105
8.1407 | -103.4126
8.1444 | -103.4133
8.1481 | -103.4121
8.1519 | -103.4123
8.1556 | -103.4128
8.1593 | -103.4141
8.163 | -103.4151
8.1667 | -103.4142
8.1704 | -103.4131
8.1741 | -103.4071
8.1778 | -103.3985
8.1815 | -103.3875
8.1852 | -103.3785
8.1889 | -103.3732
8.1926 | -103.373
8.1963 | -103.3771
8.2 | -103.3784
8.2037 | -103.3779
8.2074 | -103.3769
8.2111 | -103.3794
8.2148 | -103.3783
8.2185 | -103.3709
8.2222 | -103.3577
8.2259 | -103.3458
8.2296 | -103.3363
8.2333 | -103.3333
8.237 | -103.3302
8.2407 | -103.33
8.2444 | -103.331
8.2481 | -103.3345
8.2519 | -103.3394
8.2556 | -103.3466
8.2593 | -103.3531
8.263 | -103.3581
8.2667 | -103.3585
8.2704 | -103.3592
8.2741 | -103.3545
8.2778 | -103.3462
8.2815 | -103.3357
8.2852 | -103.322
8.2889 | -103.3021
8.2926 | -103.2782
8.2963 | -103.252
8.3 | -103.2214
8.3037 | -103.1936
8.3074 | -103.1635
8.3111 | -103.1389
8.3148 | -103.1123
8.3185 | -103.0837
8.3222 | -103.0561
8.3259 | -103.0377
8.3296 | -103.0275
8.3333 | -103.0189
8.337 | -103.0161
8.3407 | -103.0217
8.3444 | -103.0333
8.3481 | -103.0457
8.3519 | -103.0603
8.3556 | -103.0794
8.3593 | -103.0986
8.363 | -103.1185
8.3667 | -103.138
8.3704 | -103.1627
8.3741 | -103.1889
8.3778 | -103.2138
8.3815 | -103.2392
8.3852 | -103.2649
8.3889 | -103.2882
8.3926 | -103.3033
8.3963 | -103.32
8.4 | -103.3361
8.4037 | -103.3467
8.4074 | -103.3512
8.4111 | -103.3575
8.4148 | -103.3655
8.4185 | -103.3757
8.4222 | -103.3873
8.4259 | -103.392
8.4296 | -103.3973
8.4333 | -103.3955
8.437 | -103.3924
8.4407 | -103.3867
8.4444 | -103.3892
8.4481 | -103.3895
8.4519 | -103.3868
8.4556 | -103.3817
8.4593 | -103.3734
8.463 | -103.3659
8.4667 | -103.3559
8.4704 | -103.3423
8.4741 | -103.3257
8.4778 | -103.3078
8.4815 | -103.2911
8.4852 | -103.2709
8.4889 | -103.2542
8.4926 | -103.2298
8.4963 | -103.1988
8.5 | -103.1706
8.5037 | -103.1447
8.5074 | -103.1207
8.5111 | -103.0997
8.5148 | -103.079
8.5185 | -103.0554
8.5222 | -103.0346
8.5259 | -103.0184
8.5296 | -103.009
8.5333 | -102.997
8.537 | -102.9834
8.5407 | -102.9728
8.5444 | -102.9688
8.5481 | -102.9828
8.5519 | -102.9894
8.5556 | -103.0007
8.5593 | -103.0207
8.563 | -103.0466
8.5667 | -103.075
8.5704 | -103.101
8.5741 | -103.1305
8.5778 | -103.165
8.5815 | -103.2055
8.5852 | -103.2485
8.5889 | -103.2889
8.5926 | -103.331
8.5963 | -103.3726
8.6 | -103.4137
8.6037 | -103.4534
8.6074 | -103.4885
8.6111 | -103.5237
8.6148 | -103.5539
8.6185 | -103.5856
8.6222 | -103.6202
8.6259 | -103.6505
8.6296 | -103.6786
8.6333 | -103.6996
8.637 | -103.7161
8.6407 | -103.7312
8.6444 | -103.746
8.6481 | -103.7577
8.6519 | -103.7693
8.6556 | -103.7749
8.6593 | -103.7817
8.663 | -103.7854
8.6667 | -103.7844
8.6704 | -103.7744
8.6741 | -103.759
8.6778 | -103.746
8.6815 | -103.728
8.6852 | -103.7044
8.6889 | -103.6733
8.6926 | -103.6407
8.6963 | -103.6087
8.7 | -103.5713
8.7037 | -103.5259
8.7074 | -103.4807
8.7111 | -103.4319
8.7148 | -103.3833
8.7185 | -103.3398
8.7222 | -103.2985
8.7259 | -103.262
8.7296 | -103.2242
8.7333 | -103.1912
8.737 | -103.1663
8.7407 | -103.1428
8.7444 | -103.1212
8.7481 | -103.1055
8.7519 | -103.0943
8.7556 | -103.0828
8.7593 | -103.0742
8.763 | -103.075
8.7667 | -103.0804
8.7704 | -103.0891
8.7741 | -103.0959
8.7778 | -103.1005
8.7815 | -103.1082
8.7852 | -103.1234
8.7889 | -103.1414
8.7926 | -103.164
8.7963 | -103.1899
8.8 | -103.2225
8.8037 | -103.2563
8.8074 | -103.2854
8.8111 | -103.3138
8.8148 | -103.3439
8.8185 | -103.3726
8.8222 | -103.3968
8.8259 | -103.4218
8.8296 | -103.4441
8.8333 | -103.4638
8.837 | -103.4762
8.8407 | -103.4807
8.8444 | -103.4767
8.8481 | -103.4715
8.8519 | -103.4632
8.8556 | -103.4507
8.8593 | -103.4397
8.863 | -103.4248
8.8667 | -103.4118
8.8704 | -103.4023
8.8741 | -103.3939
8.8778 | -103.3861
8.8815 | -103.3834
8.8852 | -103.3879
8.8889 | -103.3977
8.8926 | -103.4074
8.8963 | -103.4191
8.9 | -103.4342
8.9037 | -103.4509
8.9074 | -103.4681
8.9111 | -103.4789
8.9148 | -103.4912
8.9185 | -103.4998
8.9222 | -103.5055
8.9259 | -103.5056
8.9296 | -103.5018
8.9333 | -103.4931
8.937 | -103.4823
8.9407 | -103.4673
8.9444 | -103.4502
8.9481 | -103.4334
8.9519 | -103.417
8.9556 | -103.3979
8.9593 | -103.3854
8.963 | -103.3801
8.9667 | -103.3797
8.9704 | -103.3886
8.9741 | -103.4058
8.9778 | -103.4235
8.9815 | -103.4464
8.9852 | -103.4787
8.9889 | -103.5149
8.9926 | -103.5545
8.9963 | -103.5985
9 | -103.6481
9.0037 | -103.6991
9.0074 | -103.7463
9.0111 | -103.7915
9.0148 | -103.8342
9.0185 | -103.8717
9.0222 | -103.9063
9.0259 | -103.9392
9.0296 | -103.9662
9.0333 | -103.9859
9.037 | -104.005
9.0407 | -104.0201
9.0444 | -104.0228
9.0481 | -104.0169
9.0519 | -104.0048
9.0556 | -103.9941
9.0593 | -103.9829
9.063 | -103.9737
9.0667 | -103.9599
9.0704 | -103.9458
9.0741 | -103.924
9.0778 | -103.899
9.0815 | -103.8645
9.0852 | -103.8351
9.0889 | -103.803
9.0926 | -103.7685
9.0963 | -103.7324
9.1 | -103.6986
9.1037 | -103.6671
9.1074 | -103.6392
9.1111 | -103.6138
9.1148 | -103.5934
9.1185 | -103.5799
9.1222 | -103.5671
9.1259 | -103.5635
9.1296 | -103.5649
9.1333 | -103.5704
9.137 | -103.5781
9.1407 | -103.5941
9.1444 | -103.6138
9.1481 | -103.6315
9.1519 | -103.656
9.1556 | -103.682
9.1593 | -103.7075
9.163 | -103.7264
9.1667 | -103.7467
9.1704 | -103.7696
9.1741 | -103.7915
9.1778 | -103.8112
9.1815 | -103.8341
9.1852 | -103.8591
9.1889 | -103.8781
9.1926 | -103.8922
9.1963 | -103.9022
9.2 | -103.9105
9.2037 | -103.917
9.2074 | -103.9239
9.2111 | -103.928
9.2148 | -103.9337
9.2185 | -103.9385
9.2222 | -103.9434
9.2259 | -103.9468
9.2296 | -103.9488
9.2333 | -103.9542
9.237 | -103.9582
9.2407 | -103.958
9.2444 | -103.9586
9.2481 | -103.9614
9.2519 | -103.9573
9.2556 | -103.9499
9.2593 | -103.9478
9.263 | -103.9401
9.2667 | -103.9309
9.2704 | -103.9226
9.2741 | -103.9184
9.2778 | -103.9063
9.2815 | -103.8912
9.2852 | -103.8807
9.2889 | -103.8661
9.2926 | -103.8534
9.2963 | -103.8409
9.3 | -103.8368
9.3037 | -103.8377
9.3074 | -103.8302
9.3111 | -103.8284
9.3148 | -103.8335
9.3185 | -103.8363
9.3222 | -103.8375
9.3259 | -103.841
9.3296 | -103.8426
9.3333 | -103.8456
9.337 | -103.8528
9.3407 | -103.8559
9.3444 | -103.8564
9.3481 | -103.8567
9.3519 | -103.8529
9.3556 | -103.8479
9.3593 | -103.8437
9.363 | -103.8388
9.3667 | -103.8335
9.3704 | -103.8279
9.3741 | -103.8274
9.3778 | -103.8292
9.3815 | -103.8314
9.3852 | -103.8357
9.3889 | -103.8431
9.3926 | -103.8475
9.3963 | -103.8513
9.4 | -103.8551
9.4037 | -103.8581
9.4074 | -103.8608
9.4111 | -103.8629
9.4148 | -103.8647
9.4185 | -103.8678
9.4222 | -103.8648
9.4259 | -103.8592
9.4296 | -103.8526
9.4333 | -103.8474
9.437 | -103.8406
9.4407 | -103.8294
9.4444 | -103.8236
9.4481 | -103.8146
9.4519 | -103.8033
9.4556 | -103.7947
9.4593 | -103.7848
9.463 | -103.7749
9.4667 | -103.7666
9.4704 | -103.757
9.4741 | -103.745
9.4778 | -103.739
9.4815 | -103.731
9.4852 | -103.7241
9.4889 | -103.717
9.4926 | -103.7111
9.4963 | -103.707
9.5 | -103.7023
9.5037 | -103.6984
9.5074 | -103.6956
9.5111 | -103.6963
9.5148 | -103.7008
9.5185 | -103.7052
9.5222 | -103.712
9.5259 | -103.7233
9.5296 | -103.7384
9.5333 | -103.7555
9.537 | -103.7747
9.5407 | -103.796
9.5444 | -103.8229
9.5481 | -103.8541
9.5519 | -103.8881
9.5556 | -103.9242
9.5593 | -103.9641
9.563 | -104.0013
9.5667 | -104.0364
9.5704 | -104.0661
9.5741 | -104.0949
9.5778 | -104.1285
9.5815 | -104.1591
9.5852 | -104.1843
9.5889 | -104.2057
9.5926 | -104.2219
9.5963 | -104.2378
9.6 | -104.2504
9.6037 | -104.2579
9.6074 | -104.2609
9.6111 | -104.2593
9.6148 | -104.2384
9.6185 | -104.2265
9.6222 | -104.2138
9.6259 | -104.1984
9.6296 | -104.1826
9.6333 | -104.1656
9.637 | -104.1472
9.6407 | -104.126
9.6444 | -104.1034
9.6481 | -104.0804
9.6519 | -104.0566
9.6556 | -104.0373
9.6593 | -104.0224
9.663 | -104.0103
9.6667 | -104.0003
9.6704 | -103.9927
9.6741 | -103.9897
9.6778 | -103.9888
9.6815 | -103.9878
9.6852 | -103.9875
9.6889 | -103.9901
9.6926 | -103.9965
9.6963 | -104.0026
9.7 | -104.0089
9.7037 | -104.0187
9.7074 | -104.0302
9.7111 | -104.0417
9.7148 | -104.0481
9.7185 | -104.0496
9.7222 | -104.0539
9.7259 | -104.0621
9.7296 | -104.0707
9.7333 | -104.0797
9.737 | -104.087
9.7407 | -104.0955
9.7444 | -104.1047
9.7481 | -104.1145
9.7519 | -104.123
9.7556 | -104.1296
9.7593 | -104.1405
9.763 | -104.1469
9.7667 | -104.1533
9.7704 | -104.1576
9.7741 | -104.1627
9.7778 | -104.17
9.7815 | -104.1769
9.7852 | -104.1862
9.7889 | -104.1944
9.7926 | -104.2027
9.7963 | -104.2092
9.8 | -104.2113
9.8037 | -104.2146
9.8074 | -104.2182
9.8111 | -104.2217
9.8148 | -104.2254
9.8185 | -104.2285
9.8222 | -104.2315
9.8259 | -104.2398
9.8296 | -104.2455
9.8333 | -104.2485
9.837 | -104.2539
9.8407 | -104.2616
9.8444 | -104.2726
9.8481 | -104.2839
9.8519 | -104.2938
9.8556 | -104.3095
9.8593 | -104.328
9.863 | -104.3497
9.8667 | -104.3708
9.8704 | -104.3886
9.8741 | -104.4024
9.8778 | -104.4213
9.8815 | -104.4403
9.8852 | -104.4523
9.8889 | -104.4606
9.8926 | -104.4648
9.8963 | -104.4668
9.9 | -104.4679
9.9037 | -104.4637
9.9074 | -104.4526
9.9111 | -104.4346
9.9148 | -104.4144
9.9185 | -104.3887
9.9222 | -104.356
9.9259 | -104.3228
9.9296 | -104.2903
9.9333 | -104.2562
9.937 | -104.2206
9.9407 | -104.1861
9.9444 | -104.156
9.9481 | -104.125
9.9519 | -104.0949
9.9556 | -104.0707
9.9593 | -104.0501
9.963 | -104.0315
9.9667 | -104.0167
9.9704 | -104.0056
9.9741 | -104.0024
9.9778 | -104.001
9.9815 | -103.9957
9.9852 | -103.9941
9.9889 | -103.9981
9.9926 | -104.0017
9.9963 | -104.0086
10 | -104.0132
10.0037 | -104.0166
10.0074 | -104.0191
10.0111 | -104.0227
10.0148 | -104.0289
10.0185 | -104.0341
10.0222 | -104.0329
10.0259 | -104.0324
10.0296 | -104.0349
10.0333 | -104.039
10.037 | -104.0433
10.0407 | -104.0483
10.0444 | -104.0528
10.0481 | -104.0571
10.0519 | -104.0636
10.0556 | -104.0712
10.0593 | -104.0758
10.063 | -104.0823
10.0667 | -104.0958
10.0704 | -104.1103
10.0741 | -104.1256
10.0778 | -104.1456
10.0815 | -104.167
10.0852 | -104.1952
10.0889 | -104.2262
10.0926 | -104.2543
10.0963 | -104.2851
10.1 | -104.3197
10.1037 | -104.3543
10.1074 | -104.3902
10.1111 | -104.4266
10.1148 | -104.4605
10.1185 | -104.4883
10.1222 | -104.516
10.1259 | -104.5487
10.1296 | -104.5713
10.1333 | -104.588
10.137 | -104.6104
10.1407 | -104.6284
10.1444 | -104.6438
10.1481 | -104.662
10.1519 | -104.6767
10.1556 | -104.6935
10.1593 | -104.7097
10.163 | -104.7214
10.1667 | -104.7347
10.1704 | -104.7471
10.1741 | -104.7565
10.1778 | -104.7663
10.1815 | -104.7759
10.1852 | -104.7789
10.1889 | -104.7796
10.1926 | -104.7796
10.1963 | -104.7792
10.2 | -104.7773
10.2037 | -104.7721
10.2074 | -104.7654
10.2111 | -104.7547
10.2148 | -104.7372
10.2185 | -104.7122
10.2222 | -104.6867
10.2259 | -104.6583
10.2296 | -104.6298
10.2333 | -104.6009
10.237 | -104.5729
10.2407 | -104.548
10.2444 | -104.5226
10.2481 | -104.4945
10.2519 | -104.4697
10.2556 | -104.4515
10.2593 | -104.4389
10.263 | -104.4277
10.2667 | -104.4246
10.2704 | -104.4273
10.2741 | -104.4283
10.2778 | -104.4317
10.2815 | -104.4384
10.2852 | -104.4469
10.2889 | -104.4561
10.2926 | -104.4676
10.2963 | -104.4785
10.3 | -104.4914
10.3037 | -104.5038
10.3074 | -104.5164
10.3111 | -104.5309
10.3148 | -104.548
10.3185 | -104.5614
10.3222 | -104.5773
10.3259 | -104.5945
10.3296 | -104.6111
10.3333 | -104.63
10.337 | -104.6524
10.3407 | -104.6735
10.3444 | -104.6949
10.3481 | -104.7147
10.3519 | -104.7334
10.3556 | -104.7465
10.3593 | -104.7585
10.363 | -104.7725
10.3667 | -104.7895
10.3704 | -104.802
10.3741 | -104.8035
10.3778 | -104.8047
10.3815 | -104.8019
10.3852 | -104.7928
10.3889 | -104.7834
10.3926 | -104.773
10.3963 | -104.7584
10.4 | -104.7455
10.4037 | -104.7307
10.4074 | -104.7115
10.4111 | -104.6923
10.4148 | -104.6763
10.4185 | -104.6636
10.4222 | -104.6526
10.4259 | -104.6467
10.4296 | -104.6455
10.4333 | -104.6441
10.437 | -104.6448
10.4407 | -104.6502
10.4444 | -104.6571
10.4481 | -104.6668
10.4519 | -104.679
10.4556 | -104.6933
10.4593 | -104.7081
10.463 | -104.7181
10.4667 | -104.727
10.4704 | -104.7373
10.4741 | -104.7432
10.4778 | -104.7448
10.4815 | -104.7424
10.4852 | -104.733
10.4889 | -104.7227
10.4926 | -104.7104
10.4963 | -104.6949
10.5 | -104.6767
10.5037 | -104.6591
10.5074 | -104.6435
10.5111 | -104.6257
10.5148 | -104.6029
10.5185 | -104.5805
10.5222 | -104.563
10.5259 | -104.5498
10.5296 | -104.5367
10.5333 | -104.5286
10.537 | -104.5274
10.5407 | -104.5263
10.5444 | -104.526
10.5481 | -104.5282
10.5519 | -104.5336
10.5556 | -104.5437
10.5593 | -104.557
10.563 | -104.569
10.5667 | -104.5814
10.5704 | -104.5906
10.5741 | -104.5956
10.5778 | -104.5992
10.5815 | -104.605
10.5852 | -104.6087
10.5889 | -104.6104
10.5926 | -104.6107
10.5963 | -104.6112
10.6 | -104.6099
10.6037 | -104.6058
10.6074 | -104.6011
10.6111 | -104.5945
10.6148 | -104.5917
10.6185 | -104.5893
10.6222 | -104.5844
10.6259 | -104.5801
10.6296 | -104.578
10.6333 | -104.5751
10.637 | -104.5755
10.6407 | -104.5758
10.6444 | -104.575
10.6481 | -104.5775
10.6519 | -104.5864
10.6556 | -104.5954
10.6593 | -104.6028
10.663 | -104.6084
10.6667 | -104.6155
10.6704 | -104.6203
10.6741 | -104.6298
10.6778 | -104.6411
10.6815 | -104.649
10.6852 | -104.648
10.6889 | -104.6485
10.6926 | -104.647
10.6963 | -104.6471
10.7 | -104.6503
10.7037 | -104.6517
10.7074 | -104.6455
10.7111 | -104.6423
10.7148 | -104.6365
10.7185 | -104.6281
10.7222 | -104.6211
10.7259 | -104.6175
10.7296 | -104.6159
10.7333 | -104.6132
10.737 | -104.6093
10.7407 | -104.6058
10.7444 | -104.6022
10.7481 | -104.5998
10.7519 | -104.5935
10.7556 | -104.5879
10.7593 | -104.5848
10.763 | -104.5861
10.7667 | -104.5881
10.7704 | -104.5874
10.7741 | -104.5877
10.7778 | -104.5873
10.7815 | -104.5842
10.7852 | -104.5835
10.7889 | -104.5864
10.7926 | -104.5872
10.7963 | -104.5869
10.8 | -104.59
10.8037 | -104.5912
10.8074 | -104.5919
10.8111 | -104.5952
10.8148 | -104.6047
10.8185 | -104.6163
10.8222 | -104.6288
10.8259 | -104.6456
10.8296 | -104.6613
10.8333 | -104.6773
10.837 | -104.693
10.8407 | -104.7097
10.8444 | -104.7299
10.8481 | -104.7506
10.8519 | -104.7706
10.8556 | -104.7877
10.8593 | -104.8067
10.863 | -104.8231
10.8667 | -104.839
10.8704 | -104.8491
10.8741 | -104.8566
10.8778 | -104.8639
10.8815 | -104.8642
10.8852 | -104.8602
10.8889 | -104.8556
10.8926 | -104.8479
10.8963 | -104.8381
10.9 | -104.8274
10.9037 | -104.8193
10.9074 | -104.8129
10.9111 | -104.8088
10.9148 | -104.7993
10.9185 | -104.7899
10.9222 | -104.783
10.9259 | -104.7823
10.9296 | -104.7872
10.9333 | -104.7935
10.937 | -104.8038
10.9407 | -104.8184
10.9444 | -104.831
10.9481 | -104.8432
10.9519 | -104.8573
10.9556 | -104.8722
10.9593 | -104.8861
10.963 | -104.9011
10.9667 | -104.9184
10.9704 | -104.9328
10.9741 | -104.9433
10.9778 | -104.9525
10.9815 | -104.9645
10.9852 | -104.9765
10.9889 | -104.9861
10.9926 | -104.9909
10.9963 | -104.9993
11 | -105.0095
11.0037 | -105.0136
11.0074 | -105.0142
11.0111 | -105.0124
11.0148 | -105.0115
11.0185 | -105.0083
11.0222 | -105.0059
11.0259 | -105.0033
11.0296 | -104.9989
11.0333 | -104.9937
11.037 | -104.9851
11.0407 | -104.9769
11.0444 | -104.9633
11.0481 | -104.9492
11.0519 | -104.9345
11.0556 | -104.9164
11.0593 | -104.8957
11.063 | -104.8712
11.0667 | -104.8462
11.0704 | -104.823
11.0741 | -104.804
11.0778 | -104.7862
11.0815 | -104.7657
11.0852 | -104.7418
11.0889 | -104.721
11.0926 | -104.7053
11.0963 | -104.6962
11.1 | -104.6907
11.1037 | -104.6903
11.1074 | -104.6921
11.1111 | -104.6958
11.1148 | -104.7037
11.1185 | -104.7125
11.1222 | -104.7244
11.1259 | -104.7365
11.1296 | -104.7536
11.1333 | -104.7747
11.137 | -104.7975
11.1407 | -104.8172
11.1444 | -104.8406
11.1481 | -104.8711
11.1519 | -104.9052
11.1556 | -104.939
11.1593 | -104.9727
11.163 | -105.0096
11.1667 | -105.0496
11.1704 | -105.0883
11.1741 | -105.1256
11.1778 | -105.1611
11.1815 | -105.195
11.1852 | -105.2279
11.1889 | -105.2601
11.1926 | -105.2895
11.1963 | -105.3157
11.2 | -105.3378
11.2037 | -105.3573
11.2074 | -105.3681
11.2111 | -105.3759
11.2148 | -105.3818
11.2185 | -105.3721
11.2222 | -105.3721
11.2259 | -105.3709
11.2296 | -105.3671
11.2333 | -105.3622
11.237 | -105.3556
11.2407 | -105.3466
11.2444 | -105.3386
11.2481 | -105.3329
11.2519 | -105.3215
11.2556 | -105.3054
11.2593 | -105.2855
11.263 | -105.2673
11.2667 | -105.2498
11.2704 | -105.2336
11.2741 | -105.2167
11.2778 | -105.2026
11.2815 | -105.191
11.2852 | -105.1797
11.2889 | -105.1669
11.2926 | -105.1536
11.2963 | -105.1446
11.3 | -105.1393
11.3037 | -105.1387
11.3074 | -105.1366
11.3111 | -105.1327
11.3148 | -105.1294
11.3185 | -105.126
11.3222 | -105.1227
11.3259 | -105.1224
11.3296 | -105.1211
11.3333 | -105.1213
11.337 | -105.1227
11.3407 | -105.1255
11.3444 | -105.1281
11.3481 | -105.1331
11.3519 | -105.1389
11.3556 | -105.1477
11.3593 | -105.1586
11.363 | -105.169
11.3667 | -105.1832
11.3704 | -105.2025
11.3741 | -105.2247
11.3778 | -105.2456
11.3815 | -105.2646
11.3852 | -105.2819
11.3889 | -105.3006
11.3926 | -105.3193
11.3963 | -105.336
11.4 | -105.3516
11.4037 | -105.3665
11.4074 | -105.3786
11.4111 | -105.3902
11.4148 | -105.3976
11.4185 | -105.4004
11.4222 | -105.4031
11.4259 | -105.3988
11.4296 | -105.3877
11.4333 | -105.3762
11.437 | -105.364
11.4407 | -105.3484
11.4444 | -105.3282
11.4481 | -105.3065
11.4519 | -105.2833
11.4556 | -105.2596
11.4593 | -105.2388
11.463 | -105.2148
11.4667 | -105.1884
11.4704 | -105.1637
11.4741 | -105.1423
11.4778 | -105.1254
11.4815 | -105.1128
11.4852 | -105.1046
11.4889 | -105.1008
11.4926 | -105.0965
11.4963 | -105.093
11.5 | -105.0949
11.5037 | -105.1002
11.5074 | -105.1062
11.5111 | -105.117
11.5148 | -105.1296
11.5185 | -105.1447
11.5222 | -105.1635
11.5259 | -105.1813
11.5296 | -105.2017
11.5333 | -105.2238
11.537 | -105.244
11.5407 | -105.2611
11.5444 | -105.2792
11.5481 | -105.2963
11.5519 | -105.3107
11.5556 | -105.321
11.5593 | -105.3219
11.563 | -105.3203
11.5667 | -105.3164
11.5704 | -105.3083
11.5741 | -105.2937
11.5778 | -105.278
11.5815 | -105.2666
11.5852 | -105.2529
11.5889 | -105.2342
11.5926 | -105.2158
11.5963 | -105.1967
11.6 | -105.1733
11.6037 | -105.1484
11.6074 | -105.1267
11.6111 | -105.1053
11.6148 | -105.0836
11.6185 | -105.0655
11.6222 | -105.0507
11.6259 | -105.0351
11.6296 | -105.0218
11.6333 | -105.0153
11.637 | -105.0113
11.6407 | -105.0067
11.6444 | -105.0072
11.6481 | -105.0121
11.6519 | -105.0176
11.6556 | -105.0225
11.6593 | -105.0319
11.663 | -105.0496
11.6667 | -105.0687
11.6704 | -105.0872
11.6741 | -105.1111
11.6778 | -105.1335
11.6815 | -105.1561
11.6852 | -105.1804
11.6889 | -105.2007
11.6926 | -105.2201
11.6963 | -105.2382
11.7 | -105.2488
11.7037 | -105.2585
11.7074 | -105.2676
11.7111 | -105.2719
11.7148 | -105.271
11.7185 | -105.2654
11.7222 | -105.2547
11.7259 | -105.2457
11.7296 | -105.2311
11.7333 | -105.2107
11.737 | -105.1855
11.7407 | -105.1641
11.7444 | -105.1381
11.7481 | -105.1064
11.7519 | -105.0849
11.7556 | -105.0611
11.7593 | -105.0378
11.763 | -105.0151
11.7667 | -104.9933
11.7704 | -104.9767
11.7741 | -104.9638
11.7778 | -104.952
11.7815 | -104.943
11.7852 | -104.9374
11.7889 | -104.9368
11.7926 | -104.9383
11.7963 | -104.9429
11.8 | -104.9512
11.8037 | -104.9585
11.8074 | -104.9692
11.8111 | -104.987
11.8148 | -105.0047
11.8185 | -105.0233
11.8222 | -105.0436
11.8259 | -105.0681
11.8296 | -105.0925
11.8333 | -105.1141
11.837 | -105.135
11.8407 | -105.1557
11.8444 | -105.1728
11.8481 | -105.1932
11.8519 | -105.2112
11.8556 | -105.2278
11.8593 | -105.245
11.863 | -105.263
11.8667 | -105.2768
11.8704 | -105.2904
11.8741 | -105.3033
11.8778 | -105.3141
11.8815 | -105.3254
11.8852 | -105.335
11.8889 | -105.3476
11.8926 | -105.3636
11.8963 | -105.3764
11.9 | -105.39
11.9037 | -105.4038
11.9074 | -105.4226
11.9111 | -105.4383
11.9148 | -105.4521
11.9185 | -105.4674
11.9222 | -105.4768
11.9259 | -105.4888
11.9296 | -105.5024
11.9333 | -105.5145
11.937 | -105.5253
11.9407 | -105.5345
11.9444 | -105.5381
11.9481 | -105.5406
11.9519 | -105.5418
11.9556 | -105.5415
11.9593 | -105.5375
11.963 | -105.5283
11.9667 | -105.5208
11.9704 | -105.5118
11.9741 | -105.5018
11.9778 | -105.4941
11.9815 | -105.4797
11.9852 | -105.4629
11.9889 | -105.4449
11.9926 | -105.4277
11.9963 | -105.4112
12 | -105.3976
12.0037 | -105.3824
12.0074 | -105.3666
12.0111 | -105.3528
12.0148 | -105.3394
12.0185 | -105.3225
12.0222 | -105.3047
12.0259 | -105.2889
12.0296 | -105.2747
12.0333 | -105.261
12.037 | -105.2452
12.0407 | -105.2304
12.0444 | -105.2188
12.0481 | -105.203
12.0519 | -105.1865
12.0556 | -105.1721
12.0593 | -105.1641
12.063 | -105.1528
12.0667 | -105.1356
12.0704 | -105.1213
12.0741 | -105.1125
12.0778 | -105.1084
12.0815 | -105.1016
12.0852 | -105.0975
12.0889 | -105.0989
12.0926 | -105.1012
12.0963 | -105.1056
12.1 | -105.1096
12.1037 | -105.1144
12.1074 | -105.1225
12.1111 | -105.1336
12.1148 | -105.1451
12.1185 | -105.1587
12.1222 | -105.172
12.1259 | -105.1852
12.1296 | -105.1998
12.1333 | -105.2162
12.137 | -105.2309
12.1407 | -105.2441
12.1444 | -105.2574
12.1481 | -105.2686
12.1519 | -105.2818
12.1556 | -105.2996
12.1593 | -105.3149
12.163 | -105.3272
12.1667 | -105.3368
12.1704 | -105.3446
12.1741 | -105.3486
12.1778 | -105.3532
12.1815 | -105.3566
12.1852 | -105.3586
12.1889 | -105.3607
12.1926 | -105.3619
12.1963 | -105.3573
12.2 | -105.3527
12.2037 | -105.3467
12.2074 | -105.339
12.2111 | -105.338
12.2148 | -105.338
12.2185 | -105.3373
12.2222 | -105.3411
12.2259 | -105.3481
12.2296 | -105.356
12.2333 | -105.3649
12.237 | -105.3768
12.2407 | -105.3939
12.2444 | -105.4139
12.2481 | -105.4313
12.2519 | -105.4453
12.2556 | -105.4619
12.2593 | -105.4803
12.263 | -105.4982
12.2667 | -105.515
12.2704 | -105.5295
12.2741 | -105.5438
12.2778 | -105.5567
12.2815 | -105.5671
12.2852 | -105.5703
12.2889 | -105.5604
12.2926 | -105.5572
12.2963 | -105.5541
12.3 | -105.5483
12.3037 | -105.5425
12.3074 | -105.5347
12.3111 | -105.5256
12.3148 | -105.5125
12.3185 | -105.4995
12.3222 | -105.487
12.3259 | -105.474
12.3296 | -105.4576
12.3333 | -105.4414
12.337 | -105.4258
12.3407 | -105.4138
12.3444 | -105.4033
12.3481 | -105.3929
12.3519 | -105.3832
12.3556 | -105.3778
12.3593 | -105.3746
12.363 | -105.3739
12.3667 | -105.3765
12.3704 | -105.3819
12.3741 | -105.3954
12.3778 | -105.4076
12.3815 | -105.4195
12.3852 | -105.4358
12.3889 | -105.4488
12.3926 | -105.4645
12.3963 | -105.4807
12.4 | -105.4959
12.4037 | -105.513
12.4074 | -105.53
12.4111 | -105.545
12.4148 | -105.5596
12.4185 | -105.5743
12.4222 | -105.5814
12.4259 | -105.5865
12.4296 | -105.5903
12.4333 | -105.5932
12.437 | -105.5895
12.4407 | -105.5808
12.4444 | -105.5723
12.4481 | -105.561
12.4519 | -105.5443
12.4556 | -105.523
12.4593 | -105.5021
12.463 | -105.4812
12.4667 | -105.4572
12.4704 | -105.4322
12.4741 | -105.4083
12.4778 | -105.3839
12.4815 | -105.3616
12.4852 | -105.3377
12.4889 | -105.3122
12.4926 | -105.29
12.4963 | -105.276
12.5 | -105.2602
12.5037 | -105.2445
12.5074 | -105.2312
12.5111 | -105.2206
12.5148 | -105.2116
12.5185 | -105.208
12.5222 | -105.2053
12.5259 | -105.2026
12.5296 | -105.2034
12.5333 | -105.2083
12.537 | -105.2128
12.5407 | -105.2179
12.5444 | -105.2228
12.5481 | -105.2307
12.5519 | -105.2416
12.5556 | -105.2504
12.5593 | -105.2623
12.563 | -105.2762
12.5667 | -105.2897
12.5704 | -105.3054
12.5741 | -105.3201
12.5778 | -105.3362
12.5815 | -105.3539
12.5852 | -105.3682
12.5889 | -105.3795
12.5926 | -105.3903
12.5963 | -105.4028
12.6 | -105.412
12.6037 | -105.4233
12.6074 | -105.4321
12.6111 | -105.4385
12.6148 | -105.4409
12.6185 | -105.4442
12.6222 | -105.4455
12.6259 | -105.4429
12.6296 | -105.4406
12.6333 | -105.4395
12.637 | -105.4371
12.6407 | -105.4325
12.6444 | -105.4277
12.6481 | -105.4222
12.6519 | -105.4181
12.6556 | -105.4127
12.6593 | -105.4036
12.663 | -105.3968
12.6667 | -105.3938
12.6704 | -105.3933
12.6741 | -105.3924
12.6778 | -105.3921
12.6815 | -105.396
12.6852 | -105.4027
12.6889 | -105.4158
12.6926 | -105.4296
12.6963 | -105.4456
12.7 | -105.4609
12.7037 | -105.4789
12.7074 | -105.4999
12.7111 | -105.5165
12.7148 | -105.5314
12.7185 | -105.547
12.7222 | -105.5659
12.7259 | -105.5862
12.7296 | -105.6007
12.7333 | -105.6122
12.737 | -105.622
12.7407 | -105.6314
12.7444 | -105.638
12.7481 | -105.6437
12.7519 | -105.6477
12.7556 | -105.649
12.7593 | -105.6471
12.763 | -105.6447
12.7667 | -105.6426
12.7704 | -105.6436
12.7741 | -105.6407
12.7778 | -105.6349
12.7815 | -105.6294
12.7852 | -105.6272
12.7889 | -105.6278
12.7926 | -105.6266
12.7963 | -105.627
12.8 | -105.6272
12.8037 | -105.6291
12.8074 | -105.6327
12.8111 | -105.6387
12.8148 | -105.6446
12.8185 | -105.6506
12.8222 | -105.6628
12.8259 | -105.6698
12.8296 | -105.6779
12.8333 | -105.6878
12.837 | -105.6966
12.8407 | -105.7057
12.8444 | -105.7159
12.8481 | -105.7241
12.8519 | -105.7291
12.8556 | -105.7348
12.8593 | -105.7377
12.863 | -105.7402
12.8667 | -105.7414
12.8704 | -105.7386
12.8741 | -105.731
12.8778 | -105.7228
12.8815 | -105.7142
12.8852 | -105.7077
12.8889 | -105.6977
12.8926 | -105.6842
12.8963 | -105.6734
12.9 | -105.661
12.9037 | -105.6482
12.9074 | -105.6319
12.9111 | -105.6174
12.9148 | -105.6041
12.9185 | -105.5889
12.9222 | -105.5754
12.9259 | -105.5627
12.9296 | -105.5533
12.9333 | -105.5463
12.937 | -105.54
12.9407 | -105.537
12.9444 | -105.5373
12.9481 | -105.5332
12.9519 | -105.5323
12.9556 | -105.5349
12.9593 | -105.544
12.963 | -105.5536
12.9667 | -105.5643
12.9704 | -105.5705
12.9741 | -105.5772
12.9778 | -105.5875
12.9815 | -105.5992
12.9852 | -105.609
12.9889 | -105.6181
12.9926 | -105.6309
12.9963 | -105.6434
13 | -105.6563
13.0037 | -105.6674
13.0074 | -105.6768
13.0111 | -105.6863
13.0148 | -105.6941
13.0185 | -105.7009
13.0222 | -105.7039
13.0259 | -105.7034
13.0296 | -105.7056
13.0333 | -105.7041
13.037 | -105.7009
13.0407 | -105.7008
13.0444 | -105.699
13.0481 | -105.6944
13.0519 | -105.688
13.0556 | -105.682
13.0593 | -105.6773
13.063 | -105.6736
13.0667 | -105.6706
13.0704 | -105.6693
13.0741 | -105.6645
13.0778 | -105.6635
13.0815 | -105.6652
13.0852 | -105.6693
13.0889 | -105.6734
13.0926 | -105.6773
13.0963 | -105.683
13.1 | -105.689
13.1037 | -105.6982
13.1074 | -105.7108
13.1111 | -105.7215
13.1148 | -105.7275
13.1185 | -105.7373
13.1222 | -105.7486
13.1259 | -105.7591
13.1296 | -105.7695
13.1333 | -105.7779
13.137 | -105.7917
13.1407 | -105.8021
13.1444 | -105.8096
13.1481 | -105.8146
13.1519 | -105.8206
13.1556 | -105.828
13.1593 | -105.8365
13.163 | -105.8436
13.1667 | -105.851
13.1704 | -105.8565
13.1741 | -105.8633
13.1778 | -105.871
13.1815 | -105.8751
13.1852 | -105.8747
13.1889 | -105.8781
13.1926 | -105.8857
13.1963 | -105.8888
13.2 | -105.89
13.2037 | -105.8911
13.2074 | -105.8941
13.2111 | -105.8994
13.2148 | -105.9039
13.2185 | -105.9015
13.2222 | -105.9028
13.2259 | -105.9023
13.2296 | -105.8956
13.2333 | -105.8892
13.237 | -105.8875
13.2407 | -105.8831
13.2444 | -105.8793
13.2481 | -105.8749
13.2519 | -105.8685
13.2556 | -105.8627
13.2593 | -105.8566
13.263 | -105.8515
13.2667 | -105.8469
13.2704 | -105.8434
13.2741 | -105.8365
13.2778 | -105.8314
13.2815 | -105.8295
13.2852 | -105.8307
13.2889 | -105.8334
13.2926 | -105.8353
13.2963 | -105.838
13.3 | -105.8474
13.3037 | -105.8628
13.3074 | -105.8755
13.3111 | -105.8879
13.3148 | -105.9044
13.3185 | -105.9233
13.3222 | -105.9447
13.3259 | -105.9632
13.3296 | -105.9807
13.3333 | -106.0018
13.337 | -106.0234
13.3407 | -106.0445
13.3444 | -106.0605
13.3481 | -106.0746
13.3519 | -106.0914
13.3556 | -106.098
13.3593 | -106.1103
13.363 | -106.1217
13.3667 | -106.131
13.3704 | -106.137
13.3741 | -106.1437
13.3778 | -106.1533
13.3815 | -106.1585
13.3852 | -106.1607
13.3889 | -106.1644
13.3926 | -106.1692
13.3963 | -106.172
13.4 | -106.1682
13.4037 | -106.1649
13.4074 | -106.1612
13.4111 | -106.1551
13.4148 | -106.1482
13.4185 | -106.1435
13.4222 | -106.1379
13.4259 | -106.1269
13.4296 | -106.1113
13.4333 | -106.0971
13.437 | -106.0822
13.4407 | -106.0674
13.4444 | -106.053
13.4481 | -106.0345
13.4519 | -106.0155
13.4556 | -105.9973
13.4593 | -105.9793
13.463 | -105.9632
13.4667 | -105.9499
13.4704 | -105.9369
13.4741 | -105.9208
13.4778 | -105.9051
13.4815 | -105.8927
13.4852 | -105.8851
13.4889 | -105.879
13.4926 | -105.8721
13.4963 | -105.8679
13.5 | -105.8641
13.5037 | -105.8652
13.5074 | -105.868
13.5111 | -105.8732
13.5148 | -105.8804
13.5185 | -105.8893
13.5222 | -105.8994
13.5259 | -105.9075
13.5296 | -105.9162
13.5333 | -105.9276
13.537 | -105.9432
13.5407 | -105.959
13.5444 | -105.9725
13.5481 | -105.9853
13.5519 | -105.9952
13.5556 | -106.0048
13.5593 | -106.0156
13.563 | -106.0247
13.5667 | -106.0326
13.5704 | -106.0404
13.5741 | -106.0448
13.5778 | -106.0476
13.5815 | -106.0481
13.5852 | -106.0468
13.5889 | -106.0427
13.5926 | -106.0399
13.5963 | -106.0359
13.6 | -106.0272
13.6037 | -106.02
13.6074 | -106.0143
13.6111 | -106.0032
13.6148 | -105.9898
13.6185 | -105.977
13.6222 | -105.9672
13.6259 | -105.9604
13.6296 | -105.9514
13.6333 | -105.9414
13.637 | -105.932
13.6407 | -105.923
13.6444 | -105.9191
13.6481 | -105.9148
13.6519 | -105.9111
13.6556 | -105.9058
13.6593 | -105.9005
13.663 | -105.8975
13.6667 | -105.8972
13.6704 | -105.8997
13.6741 | -105.8992
13.6778 | -105.8962
13.6815 | -105.8946
13.6852 | -105.8948
13.6889 | -105.8939
13.6926 | -105.8913
13.6963 | -105.8911
13.7 | -105.8906
13.7037 | -105.8895
13.7074 | -105.8898
13.7111 | -105.8893
13.7148 | -105.8883
13.7185 | -105.8901
13.7222 | -105.8914
13.7259 | -105.8926
13.7296 | -105.8913
13.7333 | -105.8915
13.737 | -105.8916
13.7407 | -105.8895
13.7444 | -105.8881
13.7481 | -105.889
13.7519 | -105.886
13.7556 | -105.882
13.7593 | -105.879
13.763 | -105.8767
13.7667 | -105.8704
13.7704 | -105.8634
13.7741 | -105.8562
13.7778 | -105.8485
13.7815 | -105.8394
13.7852 | -105.8315
13.7889 | -105.8231
13.7926 | -105.8157
13.7963 | -105.8082
13.8 | -105.7994
13.8037 | -105.7928
13.8074 | -105.7869
13.8111 | -105.782
13.8148 | -105.7768
13.8185 | -105.7701
13.8222 | -105.7603
13.8259 | -105.7534
13.8296 | -105.7478
13.8333 | -105.7434
13.837 | -105.7408
13.8407 | -105.737
13.8444 | -105.7366
13.8481 | -105.7387
13.8519 | -105.7414
13.8556 | -105.7452
13.8593 | -105.7487
13.863 | -105.7534
13.8667 | -105.7583
13.8704 | -105.7664
13.8741 | -105.7718
13.8778 | -105.7752
13.8815 | -105.7804
13.8852 | -105.7872
13.8889 | -105.7896
13.8926 | -105.7954
13.8963 | -105.8
13.9 | -105.8042
13.9037 | -105.8093
13.9074 | -105.8118
13.9111 | -105.813
13.9148 | -105.8139
13.9185 | -105.8129
13.9222 | -105.8114
13.9259 | -105.8092
13.9296 | -105.8057
13.9333 | -105.8013
13.937 | -105.798
13.9407 | -105.7952
13.9444 | -105.7921
13.9481 | -105.7877
13.9519 | -105.7829
13.9556 | -105.7763
13.9593 | -105.7671
13.963 | -105.7535
13.9667 | -105.7423
13.9704 | -105.735
13.9741 | -105.7296
13.9778 | -105.7251
13.9815 | -105.7222
13.9852 | -105.7162
13.9889 | -105.7089
13.9926 | -105.7047
13.9963 | -105.7011
14 | -105.6977
14.0037 | -105.6924
14.0074 | -105.692
14.0111 | -105.6936
14.0148 | -105.6978
14.0185 | -105.7005
14.0222 | -105.7027
14.0259 | -105.7062
14.0296 | -105.7109
14.0333 | -105.7162
14.037 | -105.7204
14.0407 | -105.7269
14.0444 | -105.7323
14.0481 | -105.7359
14.0519 | -105.742
14.0556 | -105.7483
14.0593 | -105.756
14.063 | -105.765
14.0667 | -105.7746
14.0704 | -105.7844
14.0741 | -105.7964
14.0778 | -105.8088
14.0815 | -105.8194
14.0852 | -105.8308
14.0889 | -105.8463
14.0926 | -105.8622
14.0963 | -105.8766
14.1 | -105.8922
14.1037 | -105.9048
14.1074 | -105.9182
14.1111 | -105.9343
14.1148 | -105.9509
14.1185 | -105.9657
14.1222 | -105.9776
14.1259 | -105.9882
14.1296 | -105.9982
14.1333 | -106.0088
14.137 | -106.0187
14.1407 | -106.0294
14.1444 | -106.0387
14.1481 | -106.0441
14.1519 | -106.0483
14.1556 | -106.0511
14.1593 | -106.0562
14.163 | -106.0645
14.1667 | -106.0714
14.1704 | -106.0747
14.1741 | -106.0757
14.1778 | -106.0748
14.1815 | -106.0744
14.1852 | -106.0736
14.1889 | -106.0759
14.1926 | -106.0795
14.1963 | -106.0836
14.2 | -106.0873
14.2037 | -106.0915
14.2074 | -106.0986
14.2111 | -106.0994
14.2148 | -106.0996
14.2185 | -106.1034
14.2222 | -106.1091
14.2259 | -106.1139
14.2296 | -106.1156
14.2333 | -106.1162
14.237 | -106.1187
14.2407 | -106.1228
14.2444 | -106.1262
14.2481 | -106.1283
14.2519 | -106.132
14.2556 | -106.1354
14.2593 | -106.1394
14.263 | -106.1425
14.2667 | -106.1436
14.2704 | -106.145
14.2741 | -106.1499
14.2778 | -106.1557
14.2815 | -106.1589
14.2852 | -106.1606
14.2889 | -106.1624
14.2926 | -106.1629
14.2963 | -106.1617
14.3 | -106.1612
14.3037 | -106.1624
14.3074 | -106.1671
14.3111 | -106.1713
14.3148 | -106.1718
14.3185 | -106.1712
14.3222 | -106.1741
14.3259 | -106.1771
14.3296 | -106.182
14.3333 | -106.1857
14.337 | -106.1845
14.3407 | -106.1811
14.3444 | -106.1766
14.3481 | -106.1743
14.3519 | -106.1756
14.3556 | -106.1774
14.3593 | -106.1802
14.363 | -106.1774
14.3667 | -106.1757
14.3704 | -106.1731
14.3741 | -106.1703
14.3778 | -106.1678
14.3815 | -106.1644
14.3852 | -106.1627
14.3889 | -106.1607
14.3926 | -106.1582
14.3963 | -106.1559
14.4 | -106.1563
14.4037 | -106.1583
14.4074 | -106.1593
14.4111 | -106.1579
14.4148 | -106.1623
14.4185 | -106.169
14.4222 | -106.1756
14.4259 | -106.1922
14.4296 | -106.202
14.4333 | -106.2147
14.437 | -106.2266
14.4407 | -106.2378
14.4444 | -106.2465
14.4481 | -106.2544
14.4519 | -106.2627
14.4556 | -106.2734
14.4593 | -106.2858
14.463 | -106.2964
14.4667 | -106.306
14.4704 | -106.3155
14.4741 | -106.3248
14.4778 | -106.3331
14.4815 | -106.3372
14.4852 | -106.3416
14.4889 | -106.3485
14.4926 | -106.351
14.4963 | -106.3505
14.5 | -106.3499
14.5037 | -106.3524
14.5074 | -106.3524
14.5111 | -106.349
14.5148 | -106.3423
14.5185 | -106.3362
14.5222 | -106.3325
14.5259 | -106.3282
14.5296 | -106.3201
14.5333 | -106.3167
14.537 | -106.3088
14.5407 | -106.301
14.5444 | -106.2976
14.5481 | -106.2932
14.5519 | -106.2902
14.5556 | -106.2834
14.5593 | -106.2757
14.563 | -106.265
14.5667 | -106.2514
14.5704 | -106.2425
14.5741 | -106.2331
14.5778 | -106.2239
14.5815 | -106.2095
14.5852 | -106.1929
14.5889 | -106.175
14.5926 | -106.1559
14.5963 | -106.1385
14.6 | -106.1196
14.6037 | -106.0991
14.6074 | -106.0762
14.6111 | -106.0549
14.6148 | -106.0366
14.6185 | -106.0207
14.6222 | -106.003
14.6259 | -105.9853
14.6296 | -105.9691
14.6333 | -105.9534
14.637 | -105.9403
14.6407 | -105.9307
14.6444 | -105.9232
14.6481 | -105.9196
14.6519 | -105.9193
14.6556 | -105.9221
14.6593 | -105.9275
14.663 | -105.9341
14.6667 | -105.9414
14.6704 | -105.9523
14.6741 | -105.9688
14.6778 | -105.9856
14.6815 | -106
14.6852 | -106.0199
14.6889 | -106.0447
14.6926 | -106.0704
14.6963 | -106.099
14.7 | -106.1255
14.7037 | -106.1537
14.7074 | -106.183
14.7111 | -106.2126
14.7148 | -106.239
14.7185 | -106.2657
14.7222 | -106.2915
14.7259 | -106.3163
14.7296 | -106.3388
14.7333 | -106.3583
14.737 | -106.3736
14.7407 | -106.3875
14.7444 | -106.394
14.7481 | -106.3995
14.7519 | -106.4026
14.7556 | -106.4003
14.7593 | -106.3966
14.763 | -106.3896
14.7667 | -106.3823
14.7704 | -106.3727
14.7741 | -106.3641
14.7778 | -106.3545
14.7815 | -106.3417
14.7852 | -106.3282
14.7889 | -106.3155
14.7926 | -106.3053
14.7963 | -106.2929
14.8 | -106.2805
14.8037 | -106.2675
14.8074 | -106.2526
14.8111 | -106.2402
14.8148 | -106.2294
14.8185 | -106.2179
14.8222 | -106.2069
14.8259 | -106.1956
14.8296 | -106.1854
14.8333 | -106.1739
14.837 | -106.1568
14.8407 | -106.1431
14.8444 | -106.1328
14.8481 | -106.1233
14.8519 | -106.1134
14.8556 | -106.1043
14.8593 | -106.0957
14.863 | -106.0891
14.8667 | -106.0818
14.8704 | -106.0723
14.8741 | -106.0641
14.8778 | -106.0573
14.8815 | -106.0477
14.8852 | -106.0375
14.8889 | -106.0277
14.8926 | -106.0228
14.8963 | -106.0156
14.9 | -106.0035
14.9037 | -105.991
14.9074 | -105.9805
14.9111 | -105.9733
14.9148 | -105.9637
14.9185 | -105.9517
14.9222 | -105.9422
14.9259 | -105.9351
14.9296 | -105.9274
14.9333 | -105.9217
14.937 | -105.9168
14.9407 | -105.9127
14.9444 | -105.9099
14.9481 | -105.9089
14.9519 | -105.909
14.9556 | -105.9108
14.9593 | -105.9248
14.963 | -105.9295
14.9667 | -105.9355
14.9704 | -105.9458
14.9741 | -105.9577
14.9778 | -105.9706
14.9815 | -105.9857
14.9852 | -105.9974
14.9889 | -106.0076
14.9926 | -106.0203
14.9963 | -106.0357
15 | -106.0516
15.0037 | -106.0665
15.0074 | -106.0826
15.0111 | -106.0982
15.0148 | -106.1135
15.0185 | -106.1283
15.0222 | -106.1439
15.0259 | -106.1554
15.0296 | -106.1657
15.0333 | -106.1763
15.037 | -106.1843
15.0407 | -106.1903
15.0444 | -106.1968
15.0481 | -106.2051
15.0519 | -106.2142
15.0556 | -106.2187
15.0593 | -106.2224
15.063 | -106.2266
15.0667 | -106.2293
15.0704 | -106.2282
15.0741 | -106.2296
15.0778 | -106.2326
15.0815 | -106.2376
15.0852 | -106.2455
15.0889 | -106.2542
15.0926 | -106.2641
15.0963 | -106.2727
15.1 | -106.2827
15.1037 | -106.2956
15.1074 | -106.3074
15.1111 | -106.3194
15.1148 | -106.3361
15.1185 | -106.3539
15.1222 | -106.3673
15.1259 | -106.379
15.1296 | -106.3923
15.1333 | -106.4072
15.137 | -106.4189
15.1407 | -106.4284
15.1444 | -106.4385
15.1481 | -106.4461
15.1519 | -106.4529
15.1556 | -106.4587
15.1593 | -106.4639
15.163 | -106.4673
15.1667 | -106.4698
15.1704 | -106.469
15.1741 | -106.4633
15.1778 | -106.4553
15.1815 | -106.4493
15.1852 | -106.4412
15.1889 | -106.4283
15.1926 | -106.4114
15.1963 | -106.3917
15.2 | -106.3728
15.2037 | -106.3604
15.2074 | -106.3476
15.2111 | -106.3294
15.2148 | -106.3065
15.2185 | -106.2853
15.2222 | -106.265
15.2259 | -106.2457
15.2296 | -106.2285
15.2333 | -106.2084
15.237 | -106.1888
15.2407 | -106.1688
15.2444 | -106.1491
15.2481 | -106.1315
15.2519 | -106.115
15.2556 | -106.1003
15.2593 | -106.0881
15.263 | -106.0765
15.2667 | -106.0647
15.2704 | -106.0577
15.2741 | -106.0524
15.2778 | -106.0473
15.2815 | -106.0434
15.2852 | -106.0407
15.2889 | -106.0413
15.2926 | -106.041
15.2963 | -106.0406
15.3 | -106.0429
15.3037 | -106.0475
15.3074 | -106.0513
15.3111 | -106.0539
15.3148 | -106.0587
15.3185 | -106.0665
15.3222 | -106.0748
15.3259 | -106.0841
15.3296 | -106.093
15.3333 | -106.099
15.337 | -106.1091
15.3407 | -106.1199
15.3444 | -106.1334
15.3481 | -106.1499
15.3519 | -106.1641
15.3556 | -106.1771
15.3593 | -106.1893
15.363 | -106.1997
15.3667 | -106.2108
15.3704 | -106.2226
15.3741 | -106.2345
15.3778 | -106.2456
15.3815 | -106.2574
15.3852 | -106.2668
15.3889 | -106.277
15.3926 | -106.2881
15.3963 | -106.2976
15.4 | -106.3055
15.4037 | -106.3126
15.4074 | -106.3198
15.4111 | -106.3274
15.4148 | -106.336
15.4185 | -106.3406
15.4222 | -106.3419
15.4259 | -106.3432
15.4296 | -106.3486
15.4333 | -106.3535
15.437 | -106.3526
15.4407 | -106.3543
15.4444 | -106.355
15.4481 | -106.3558
15.4519 | -106.3588
15.4556 | -106.3596
15.4593 | -106.3597
15.463 | -106.362
15.4667 | -106.363
15.4704 | -106.3614
15.4741 | -106.3597
15.4778 | -106.358
15.4815 | -106.3556
15.4852 | -106.3542
15.4889 | -106.3537
15.4926 | -106.3551
15.4963 | -106.354
15.5 | -106.3559
15.5037 | -106.3574
15.5074 | -106.3613
15.5111 | -106.3674
15.5148 | -106.3738
15.5185 | -106.381
15.5222 | -106.3851
15.5259 | -106.3886
15.5296 | -106.3958
15.5333 | -106.4065
15.537 | -106.42
15.5407 | -106.4308
15.5444 | -106.4413
15.5481 | -106.4524
15.5519 | -106.4635
15.5556 | -106.4792
15.5593 | -106.4921
15.563 | -106.5034
15.5667 | -106.5145
15.5704 | -106.5262
15.5741 | -106.5354
15.5778 | -106.5447
15.5815 | -106.5562
15.5852 | -106.5668
15.5889 | -106.5749
15.5926 | -106.5786
15.5963 | -106.578
15.6 | -106.5791
15.6037 | -106.58
15.6074 | -106.5791
15.6111 | -106.578
15.6148 | -106.5728
15.6185 | -106.5683
15.6222 | -106.5618
15.6259 | -106.5551
15.6296 | -106.547
15.6333 | -106.5393
15.637 | -106.5285
15.6407 | -106.5158
15.6444 | -106.5044
15.6481 | -106.4914
15.6519 | -106.473
15.6556 | -106.4569
15.6593 | -106.4433
15.663 | -106.4264
15.6667 | -106.408
15.6704 | -106.3898
15.6741 | -106.3728
15.6778 | -106.3559
15.6815 | -106.3381
15.6852 | -106.3197
15.6889 | -106.3019
15.6926 | -106.2863
15.6963 | -106.2674
15.7 | -106.2494
15.7037 | -106.233
15.7074 | -106.2185
15.7111 | -106.2059
15.7148 | -106.193
15.7185 | -106.1821
15.7222 | -106.1726
15.7259 | -106.1625
15.7296 | -106.1555
15.7333 | -106.1501
15.737 | -106.1457
15.7407 | -106.1427
15.7444 | -106.1419
15.7481 | -106.1442
15.7519 | -106.1462
15.7556 | -106.1484
15.7593 | -106.1515
15.763 | -106.1535
15.7667 | -106.1532
15.7704 | -106.1541
15.7741 | -106.1578
15.7778 | -106.163
15.7815 | -106.1647
15.7852 | -106.1675
15.7889 | -106.1732
15.7926 | -106.1788
15.7963 | -106.1823
15.8 | -106.1879
15.8037 | -106.1927
15.8074 | -106.196
15.8111 | -106.2002
15.8148 | -106.206
15.8185 | -106.2111
15.8222 | -106.2142
15.8259 | -106.2128
15.8296 | -106.2123
15.8333 | -106.2096
15.837 | -106.2035
15.8407 | -106.1997
15.8444 | -106.1964
15.8481 | -106.1929
15.8519 | -106.1857
15.8556 | -106.1791
15.8593 | -106.1719
15.863 | -106.1649
15.8667 | -106.1596
15.8704 | -106.1515
15.8741 | -106.1437
15.8778 | -106.1367
15.8815 | -106.1315
15.8852 | -106.1268
15.8889 | -106.1242
15.8926 | -106.121
15.8963 | -106.1151
15.9 | -106.111
15.9037 | -106.1091
15.9074 | -106.1095
15.9111 | -106.1109
15.9148 | -106.112
15.9185 | -106.1123
15.9222 | -106.1158
15.9259 | -106.1143
15.9296 | -106.1141
15.9333 | -106.1148
15.937 | -106.113
15.9407 | -106.1155
15.9444 | -106.1188
15.9481 | -106.1194
15.9519 | -106.1203
15.9556 | -106.1216
15.9593 | -106.1227
15.963 | -106.1229
15.9667 | -106.1234
15.9704 | -106.1213
15.9741 | -106.1228
15.9778 | -106.1264
15.9815 | -106.1277
15.9852 | -106.1281
15.9889 | -106.1297
15.9926 | -106.1372
15.9963 | -106.1446
16 | -106.1486
16.0037 | -106.1531
16.0074 | -106.1607
16.0111 | -106.1687
16.0148 | -106.1782
16.0185 | -106.1873
16.0222 | -106.1992
16.0259 | -106.2115
16.0296 | -106.2301
16.0333 | -106.2424
16.037 | -106.2564
16.0407 | -106.2716
16.0444 | -106.2855
16.0481 | -106.3018
16.0519 | -106.322
16.0556 | -106.3409
16.0593 | -106.3572
16.063 | -106.3753
16.0667 | -106.3951
16.0704 | -106.4137
16.0741 | -106.4304
16.0778 | -106.4446
16.0815 | -106.4597
16.0852 | -106.474
16.0889 | -106.4844
16.0926 | -106.4973
16.0963 | -106.5085
16.1 | -106.5199
16.1037 | -106.5306
16.1074 | -106.5412
16.1111 | -106.5477
16.1148 | -106.5524
16.1185 | -106.5551
16.1222 | -106.5585
16.1259 | -106.5608
16.1296 | -106.5578
16.1333 | -106.5547
16.137 | -106.5517
16.1407 | -106.5501
16.1444 | -106.5465
16.1481 | -106.5419
16.1519 | -106.5358
16.1556 | -106.5325
16.1593 | -106.5261
16.163 | -106.5208
16.1667 | -106.515
16.1704 | -106.5053
16.1741 | -106.4974
16.1778 | -106.4921
16.1815 | -106.4856
16.1852 | -106.4754
16.1889 | -106.4664
16.1926 | -106.4582
16.1963 | -106.4511
16.2 | -106.445
16.2037 | -106.4346
16.2074 | -106.4235
16.2111 | -106.4143
16.2148 | -106.4065
16.2185 | -106.3991
16.2222 | -106.3898
16.2259 | -106.381
16.2296 | -106.3732
16.2333 | -106.3663
16.237 | -106.3583
16.2407 | -106.3495
16.2444 | -106.3402
16.2481 | -106.3328
16.2519 | -106.3246
16.2556 | -106.3173
16.2593 | -106.3137
16.263 | -106.3099
16.2667 | -106.3072
16.2704 | -106.3038
16.2741 | -106.2992
16.2778 | -106.2973
16.2815 | -106.2968
16.2852 | -106.2969
16.2889 | -106.3003
16.2926 | -106.3047
16.2963 | -106.3091
16.3 | -106.3164
16.3037 | -106.3256
16.3074 | -106.3343
16.3111 | -106.3436
16.3148 | -106.3536
16.3185 | -106.3649
16.3222 | -106.3811
16.3259 | -106.3963
16.3296 | -106.412
16.3333 | -106.4292
16.337 | -106.447
16.3407 | -106.4655
16.3444 | -106.4851
16.3481 | -106.5056
16.3519 | -106.5282
16.3556 | -106.5507
16.3593 | -106.5747
16.363 | -106.6007
16.3667 | -106.6258
16.3704 | -106.6506
16.3741 | -106.6733
16.3778 | -106.6973
16.3815 | -106.721
16.3852 | -106.7439
16.3889 | -106.7665
16.3926 | -106.7868
16.3963 | -106.8035
16.4 | -106.8196
16.4037 | -106.8343
16.4074 | -106.8446
16.4111 | -106.8532
16.4148 | -106.8629
16.4185 | -106.8745
16.4222 | -106.8822
16.4259 | -106.8866
16.4296 | -106.8886
16.4333 | -106.8906
16.437 | -106.8911
16.4407 | -106.8875
16.4444 | -106.8802
16.4481 | -106.8713
16.4519 | -106.8599
16.4556 | -106.8461
16.4593 | -106.8346
16.463 | -106.8215
16.4667 | -106.8068
16.4704 | -106.7926
16.4741 | -106.7737
16.4778 | -106.7549
16.4815 | -106.7374
16.4852 | -106.7194
16.4889 | -106.6997
16.4926 | -106.6786
16.4963 | -106.6598
16.5 | -106.6417
16.5037 | -106.6238
16.5074 | -106.6045
16.5111 | -106.5855
16.5148 | -106.5671
16.5185 | -106.5488
16.5222 | -106.53
16.5259 | -106.5128
16.5296 | -106.4963
16.5333 | -106.4796
16.537 | -106.4646
16.5407 | -106.4504
16.5444 | -106.4387
16.5481 | -106.4308
16.5519 | -106.4221
16.5556 | -106.411
16.5593 | -106.4018
16.563 | -106.3971
16.5667 | -106.3889
16.5704 | -106.383
16.5741 | -106.3766
16.5778 | -106.3714
16.5815 | -106.3684
16.5852 | -106.3655
16.5889 | -106.3637
16.5926 | -106.3635
16.5963 | -106.366
16.6 | -106.366
16.6037 | -106.3641
16.6074 | -106.3628
16.6111 | -106.3636
16.6148 | -106.365
16.6185 | -106.3674
16.6222 | -106.3705
16.6259 | -106.3749
16.6296 | -106.3781
16.6333 | -106.3844
16.637 | -106.3914
16.6407 | -106.3975
16.6444 | -106.4032
16.6481 | -106.4081
16.6519 | -106.4109
16.6556 | -106.4144
16.6593 | -106.4184
16.663 | -106.423
16.6667 | -106.4259
16.6704 | -106.4284
16.6741 | -106.4325
16.6778 | -106.4384
16.6815 | -106.4432
16.6852 | -106.4453
16.6889 | -106.4492
16.6926 | -106.4561
16.6963 | -106.46
16.7 | -106.4637
16.7037 | -106.4661
16.7074 | -106.4702
16.7111 | -106.4789
16.7148 | -106.486
16.7185 | -106.49
16.7222 | -106.4961
16.7259 | -106.5022
16.7296 | -106.5073
16.7333 | -106.5126
16.737 | -106.5171
16.7407 | -106.5215
16.7444 | -106.5272
16.7481 | -106.5318
16.7519 | -106.5365
16.7556 | -106.5419
16.7593 | -106.5458
16.763 | -106.5504
16.7667 | -106.5567
16.7704 | -106.5594
16.7741 | -106.5617
16.7778 | -106.5674
16.7815 | -106.5712
16.7852 | -106.571
16.7889 | -106.5739
16.7926 | -106.579
16.7963 | -106.5833
16.8 | -106.5868
16.8037 | -106.5897
16.8074 | -106.5925
16.8111 | -106.595
16.8148 | -106.5974
16.8185 | -106.5995
16.8222 | -106.6015
16.8259 | -106.6042
16.8296 | -106.6058
16.8333 | -106.6064
16.837 | -106.6074
16.8407 | -106.6126
16.8444 | -106.6191
16.8481 | -106.6242
16.8519 | -106.6264
16.8556 | -106.631
16.8593 | -106.6364
16.863 | -106.6415
16.8667 | -106.6444
16.8704 | -106.6518
16.8741 | -106.6589
16.8778 | -106.6673
16.8815 | -106.6776
16.8852 | -106.6892
16.8889 | -106.7013
16.8926 | -106.715
16.8963 | -106.7277
16.9 | -106.7397
16.9037 | -106.7523
16.9074 | -106.7669
16.9111 | -106.781
16.9148 | -106.795
16.9185 | -106.8073
16.9222 | -106.82
16.9259 | -106.833
16.9296 | -106.8437
16.9333 | -106.8494
16.937 | -106.8569
16.9407 | -106.8659
16.9444 | -106.8746
16.9481 | -106.8792
16.9519 | -106.8844
16.9556 | -106.8901
16.9593 | -106.8953
16.963 | -106.8997
16.9667 | -106.9036
16.9704 | -106.904
16.9741 | -106.9018
16.9778 | -106.8986
16.9815 | -106.8969
16.9852 | -106.8964
16.9889 | -106.8947
16.9926 | -106.8891
16.9963 | -106.881
17 | -106.8761
17.0037 | -106.8714
17.0074 | -106.8636
17.0111 | -106.8567
17.0148 | -106.8493
17.0185 | -106.8406
17.0222 | -106.8342
17.0259 | -106.8277
17.0296 | -106.8163
17.0333 | -106.8031
17.037 | -106.7919
17.0407 | -106.7821
17.0444 | -106.7723
17.0481 | -106.7616
17.0519 | -106.7508
17.0556 | -106.7422
17.0593 | -106.7326
17.063 | -106.7209
17.0667 | -106.7083
17.0704 | -106.6973
17.0741 | -106.6849
17.0778 | -106.6705
17.0815 | -106.6571
17.0852 | -106.6454
17.0889 | -106.6362
17.0926 | -106.6289
17.0963 | -106.6231
17.1 | -106.6151
17.1037 | -106.6086
17.1074 | -106.6023
17.1111 | -106.5977
17.1148 | -106.5911
17.1185 | -106.5834
17.1222 | -106.5772
17.1259 | -106.574
17.1296 | -106.5717
17.1333 | -106.5695
17.137 | -106.5681
17.1407 | -106.5682
17.1444 | -106.569
17.1481 | -106.5714
17.1519 | -106.5746
17.1556 | -106.5786
17.1593 | -106.5804
17.163 | -106.5858
17.1667 | -106.593
17.1704 | -106.6023
17.1741 | -106.6133
17.1778 | -106.6264
17.1815 | -106.6373
17.1852 | -106.6469
17.1889 | -106.6555
17.1926 | -106.6665
17.1963 | -106.6761
17.2 | -106.6866
17.2037 | -106.6992
17.2074 | -106.7073
17.2111 | -106.7163
17.2148 | -106.7258
17.2185 | -106.7345
17.2222 | -106.7442
17.2259 | -106.7497
17.2296 | -106.75
17.2333 | -106.7508
17.237 | -106.7544
17.2407 | -106.7582
17.2444 | -106.7595
17.2481 | -106.7563
17.2519 | -106.7552
17.2556 | -106.7562
17.2593 | -106.753
17.263 | -106.7473
17.2667 | -106.7409
17.2704 | -106.7367
17.2741 | -106.7344
17.2778 | -106.7303
17.2815 | -106.7216
17.2852 | -106.7146
17.2889 | -106.7087
17.2926 | -106.7024
17.2963 | -106.6966
17.3 | -106.6917
17.3037 | -106.6875
17.3074 | -106.6842
17.3111 | -106.6824
17.3148 | -106.6773
17.3185 | -106.6741
17.3222 | -106.6727
17.3259 | -106.6713
17.3296 | -106.6685
17.3333 | -106.664
17.337 | -106.6605
17.3407 | -106.6561
17.3444 | -106.6529
17.3481 | -106.6522
17.3519 | -106.65
17.3556 | -106.6499
17.3593 | -106.6489
17.363 | -106.65
17.3667 | -106.6523
17.3704 | -106.6533
17.3741 | -106.6538
17.3778 | -106.6539
17.3815 | -106.6524
17.3852 | -106.6532
17.3889 | -106.6561
17.3926 | -106.659
17.3963 | -106.6611
17.4 | -106.6636
17.4037 | -106.6667
17.4074 | -106.6723
17.4111 | -106.6763
17.4148 | -106.6811
17.4185 | -106.6823
17.4222 | -106.6811
17.4259 | -106.6803
17.4296 | -106.6801
17.4333 | -106.6824
17.437 | -106.6877
17.4407 | -106.6886
17.4444 | -106.6883
17.4481 | -106.6889
17.4519 | -106.6886
17.4556 | -106.6913
17.4593 | -106.6928
17.463 | -106.6942
17.4667 | -106.6982
17.4704 | -106.698
17.4741 | -106.6993
17.4778 | -106.7025
17.4815 | -106.7042
17.4852 | -106.7049
17.4889 | -106.7057
17.4926 | -106.7059
17.4963 | -106.7072
17.5 | -106.7075
17.5037 | -106.709
17.5074 | -106.7149
17.5111 | -106.7217
17.5148 | -106.729
17.5185 | -106.7348
17.5222 | -106.7377
17.5259 | -106.7443
17.5296 | -106.7548
17.5333 | -106.7621
17.537 | -106.7681
17.5407 | -106.7762
17.5444 | -106.7838
17.5481 | -106.79
17.5519 | -106.7979
17.5556 | -106.8063
17.5593 | -106.815
17.563 | -106.8254
17.5667 | -106.8362
17.5704 | -106.8448
17.5741 | -106.852
17.5778 | -106.8601
17.5815 | -106.8713
17.5852 | -106.8824
17.5889 | -106.8905
17.5926 | -106.8954
17.5963 | -106.9013
17.6 | -106.9101
17.6037 | -106.9196
17.6074 | -106.9297
17.6111 | -106.9384
17.6148 | -106.9454
17.6185 | -106.9546
17.6222 | -106.9617
17.6259 | -106.9682
17.6296 | -106.9732
17.6333 | -106.9815
17.637 | -106.9871
17.6407 | -106.9905
17.6444 | -106.9946
17.6481 | -106.9996
17.6519 | -107.0026
17.6556 | -107.0038
17.6593 | -107.0059
17.663 | -107.0065
17.6667 | -107.0042
17.6704 | -107.0024
17.6741 | -107.0049
17.6778 | -107.0075
17.6815 | -107.0075
17.6852 | -107.0023
17.6889 | -106.9927
17.6926 | -106.9849
17.6963 | -106.9749
17.7 | -106.9649
17.7037 | -106.9576
17.7074 | -106.9481
17.7111 | -106.9388
17.7148 | -106.9294
17.7185 | -106.9166
17.7222 | -106.9077
17.7259 | -106.8992
17.7296 | -106.889
17.7333 | -106.8761
17.737 | -106.8641
17.7407 | -106.8547
17.7444 | -106.8449
17.7481 | -106.8349
17.7519 | -106.8235
17.7556 | -106.8121
17.7593 | -106.8018
17.763 | -106.7923
17.7667 | -106.7845
17.7704 | -106.7782
17.7741 | -106.7743
17.7778 | -106.7686
17.7815 | -106.7619
17.7852 | -106.7596
17.7889 | -106.7596
17.7926 | -106.7568
17.7963 | -106.7553
17.8 | -106.7558
17.8037 | -106.7559
17.8074 | -106.756
17.8111 | -106.7563
17.8148 | -106.7569
17.8185 | -106.7582
17.8222 | -106.7607
17.8259 | -106.7598
17.8296 | -106.7618
17.8333 | -106.7655
17.837 | -106.7695
17.8407 | -106.7772
17.8444 | -106.7859
17.8481 | -106.7917
17.8519 | -106.7951
17.8556 | -106.798
17.8593 | -106.8018
17.863 | -106.8075
17.8667 | -106.8136
17.8704 | -106.8182
17.8741 | -106.8242
17.8778 | -106.8302
17.8815 | -106.8328
17.8852 | -106.8373
17.8889 | -106.8417
17.8926 | -106.8434
17.8963 | -106.8457
17.9 | -106.8482
17.9037 | -106.8484
17.9074 | -106.8514
17.9111 | -106.8561
17.9148 | -106.8578
17.9185 | -106.8596
17.9222 | -106.8619
17.9259 | -106.8663
17.9296 | -106.8679
17.9333 | -106.8671
17.937 | -106.8702
17.9407 | -106.8749
17.9444 | -106.8811
17.9481 | -106.8846
17.9519 | -106.8901
17.9556 | -106.8917
17.9593 | -106.8912
17.963 | -106.8939
17.9667 | -106.8985
17.9704 | -106.9033
17.9741 | -106.9071
17.9778 | -106.9096
17.9815 | -106.9108
17.9852 | -106.9132
17.9889 | -106.918
17.9926 | -106.9213
17.9963 | -106.9207
18 | -106.9195
18.0037 | -106.9205
18.0074 | -106.922
18.0111 | -106.9209
18.0148 | -106.9172
18.0185 | -106.9139
18.0222 | -106.9084
18.0259 | -106.9022
18.0296 | -106.8996
18.0333 | -106.8982
18.037 | -106.896
18.0407 | -106.8911
18.0444 | -106.8838
18.0481 | -106.8776
18.0519 | -106.872
18.0556 | -106.8675
18.0593 | -106.8609
18.063 | -106.8565
18.0667 | -106.851
18.0704 | -106.847
18.0741 | -106.8415
18.0778 | -106.8373
18.0815 | -106.8371
18.0852 | -106.8342
18.0889 | -106.8275
18.0926 | -106.8213
18.0963 | -106.8188
18.1 | -106.8182
18.1037 | -106.8168
18.1074 | -106.8155
18.1111 | -106.8149
18.1148 | -106.8157
18.1185 | -106.8154
18.1222 | -106.8145
18.1259 | -106.8168
18.1296 | -106.8211
18.1333 | -106.8261
18.137 | -106.831
18.1407 | -106.8358
18.1444 | -106.8403
18.1481 | -106.8472
18.1519 | -106.853
18.1556 | -106.8596
18.1593 | -106.8669
18.163 | -106.873
18.1667 | -106.8813
18.1704 | -106.8861
18.1741 | -106.8933
18.1778 | -106.9017
18.1815 | -106.91
18.1852 | -106.9181
18.1889 | -106.926
18.1926 | -106.9334
18.1963 | -106.9403
18.2 | -106.9463
18.2037 | -106.9533
18.2074 | -106.9631
18.2111 | -106.9706
18.2148 | -106.9786
18.2185 | -106.9865
18.2222 | -106.9929
18.2259 | -106.9985
18.2296 | -107.0039
18.2333 | -107.0129
18.237 | -107.022
18.2407 | -107.0305
18.2444 | -107.0396
18.2481 | -107.0496
18.2519 | -107.0581
18.2556 | -107.0668
18.2593 | -107.0768
18.263 | -107.0871
18.2667 | -107.0958
18.2704 | -107.1043
18.2741 | -107.1155
18.2778 | -107.1281
18.2815 | -107.1408
18.2852 | -107.1525
18.2889 | -107.1624
18.2926 | -107.1702
18.2963 | -107.1774
18.3 | -107.1861
18.3037 | -107.1958
18.3074 | -107.2038
18.3111 | -107.2108
18.3148 | -107.2189
18.3185 | -107.2311
18.3222 | -107.2416
18.3259 | -107.2516
18.3296 | -107.2592
18.3333 | -107.2659
18.337 | -107.2714
18.3407 | -107.2743
18.3444 | -107.2765
18.3481 | -107.2817
18.3519 | -107.2883
18.3556 | -107.2913
18.3593 | -107.2929
18.363 | -107.2946
18.3667 | -107.2969
18.3704 | -107.2995
18.3741 | -107.3011
18.3778 | -107.3032
18.3815 | -107.3066
18.3852 | -107.3064
18.3889 | -107.3047
18.3926 | -107.3025
18.3963 | -107.3014
18.4 | -107.2999
18.4037 | -107.3006
18.4074 | -107.302
18.4111 | -107.3028
18.4148 | -107.3
18.4185 | -107.2962
18.4222 | -107.2921
18.4259 | -107.2893
18.4296 | -107.2869
18.4333 | -107.2848
18.437 | -107.2814
18.4407 | -107.2785
18.4444 | -107.2755
18.4481 | -107.2719
18.4519 | -107.2692
18.4556 | -107.2698
18.4593 | -107.2683
18.463 | -107.2648
18.4667 | -107.2614
18.4704 | -107.258
18.4741 | -107.2566
18.4778 | -107.2545
18.4815 | -107.2516
18.4852 | -107.2498
18.4889 | -107.2464
18.4926 | -107.2448
18.4963 | -107.2416
18.5 | -107.2366
18.5037 | -107.2318
18.5074 | -107.2273
18.5111 | -107.222
18.5148 | -107.2193
18.5185 | -107.2172
18.5222 | -107.2143
18.5259 | -107.21
18.5296 | -107.2069
18.5333 | -107.2043
18.537 | -107.1984
18.5407 | -107.1901
18.5444 | -107.1835
18.5481 | -107.1783
18.5519 | -107.1713
18.5556 | -107.1666
18.5593 | -107.1592
18.563 | -107.1507
18.5667 | -107.1454
18.5704 | -107.1388
18.5741 | -107.1314
18.5778 | -107.1212
18.5815 | -107.11
18.5852 | -107.0988
18.5889 | -107.0891
18.5926 | -107.0816
18.5963 | -107.0718
18.6 | -107.0597
18.6037 | -107.0503
18.6074 | -107.0388
18.6111 | -107.0263
18.6148 | -107.0132
18.6185 | -107.0002
18.6222 | -106.9863
18.6259 | -106.9716
18.6296 | -106.958
18.6333 | -106.9441
18.637 | -106.9299
18.6407 | -106.917
18.6444 | -106.903
18.6481 | -106.8901
18.6519 | -106.8773
18.6556 | -106.8634
18.6593 | -106.8491
18.663 | -106.8348
18.6667 | -106.8187
18.6704 | -106.8038
18.6741 | -106.7902
18.6778 | -106.7785
18.6815 | -106.7664
18.6852 | -106.7533
18.6889 | -106.7413
18.6926 | -106.7299
18.6963 | -106.7206
18.7 | -106.7154
18.7037 | -106.7085
18.7074 | -106.7015
18.7111 | -106.6937
18.7148 | -106.6874
18.7185 | -106.6818
18.7222 | -106.6764
18.7259 | -106.6731
18.7296 | -106.6733
18.7333 | -106.6718
18.737 | -106.6681
18.7407 | -106.6652
18.7444 | -106.6648
18.7481 | -106.6659
18.7519 | -106.6676
18.7556 | -106.6694
18.7593 | -106.6708
18.763 | -106.6749
18.7667 | -106.6812
18.7704 | -106.6877
18.7741 | -106.6915
18.7778 | -106.6945
18.7815 | -106.698
18.7852 | -106.7028
18.7889 | -106.7084
18.7926 | -106.7143
18.7963 | -106.7171
18.8 | -106.7218
18.8037 | -106.729
18.8074 | -106.7362
18.8111 | -106.7421
18.8148 | -106.7478
18.8185 | -106.7539
18.8222 | -106.7596
18.8259 | -106.7683
18.8296 | -106.7769
18.8333 | -106.7812
18.837 | -106.7873
18.8407 | -106.7939
18.8444 | -106.8005
18.8481 | -106.807
18.8519 | -106.814
18.8556 | -106.8203
18.8593 | -106.8252
18.863 | -106.8329
18.8667 | -106.8412
18.8704 | -106.8477
18.8741 | -106.8543
18.8778 | -106.8599
18.8815 | -106.8652
18.8852 | -106.8714
18.8889 | -106.8799
18.8926 | -106.8905
18.8963 | -106.9
18.9 | -106.9082
18.9037 | -106.9171
18.9074 | -106.9271
18.9111 | -106.9391
18.9148 | -106.9492
18.9185 | -106.9599
18.9222 | -106.9719
18.9259 | -106.9848
18.9296 | -106.998
18.9333 | -107.0135
18.937 | -107.0266
18.9407 | -107.0392
18.9444 | -107.0524
18.9481 | -107.0674
18.9519 | -107.0827
18.9556 | -107.0957
18.9593 | -107.1069
18.963 | -107.118
18.9667 | -107.1289
18.9704 | -107.1377
18.9741 | -107.146
18.9778 | -107.1528
18.9815 | -107.1591
18.9852 | -107.1619
18.9889 | -107.166
18.9926 | -107.1696
18.9963 | -107.17
19 | -107.1708
19.0037 | -107.1731
19.0074 | -107.1748
19.0111 | -107.1701
19.0148 | -107.1642
19.0185 | -107.1561
19.0222 | -107.1451
19.0259 | -107.1347
19.0296 | -107.1218
19.0333 | -107.106
19.037 | -107.0911
19.0407 | -107.0752
19.0444 | -107.0606
19.0481 | -107.0436
19.0519 | -107.0241
19.0556 | -107.0056
19.0593 | -106.9881
19.063 | -106.9675
19.0667 | -106.9514
19.0704 | -106.9323
19.0741 | -106.9144
19.0778 | -106.8984
19.0815 | -106.884
19.0852 | -106.8704
19.0889 | -106.8586
19.0926 | -106.8473
19.0963 | -106.8364
19.1 | -106.8273
19.1037 | -106.8205
19.1074 | -106.8155
19.1111 | -106.8113
19.1148 | -106.8088
19.1185 | -106.8091
19.1222 | -106.8098
19.1259 | -106.8123
19.1296 | -106.8177
19.1333 | -106.8254
19.137 | -106.8343
19.1407 | -106.8426
19.1444 | -106.8538
19.1481 | -106.8654
19.1519 | -106.8778
19.1556 | -106.8925
19.1593 | -106.9063
19.163 | -106.9187
19.1667 | -106.9302
19.1704 | -106.9421
19.1741 | -106.9539
19.1778 | -106.9663
19.1815 | -106.9792
19.1852 | -106.9929
19.1889 | -107.0074
19.1926 | -107.0218
19.1963 | -107.034
19.2 | -107.045
19.2037 | -107.0533
19.2074 | -107.0633
19.2111 | -107.0739
19.2148 | -107.0823
19.2185 | -107.0885
19.2222 | -107.0946
19.2259 | -107.0996
19.2296 | -107.1067
19.2333 | -107.1074
19.237 | -107.1109
19.2407 | -107.1151
19.2444 | -107.1201
19.2481 | -107.1234
19.2519 | -107.1266
19.2556 | -107.1311
19.2593 | -107.1362
19.263 | -107.1369
19.2667 | -107.1368
19.2704 | -107.1399
19.2741 | -107.1397
19.2778 | -107.1398
19.2815 | -107.1412
19.2852 | -107.1438
19.2889 | -107.1456
19.2926 | -107.1498
19.2963 | -107.1536
19.3 | -107.1554
19.3037 | -107.1569
19.3074 | -107.1574
19.3111 | -107.1581
19.3148 | -107.1606
19.3185 | -107.1625
19.3222 | -107.1654
19.3259 | -107.1679
19.3296 | -107.1692
19.3333 | -107.1676
19.337 | -107.1679
19.3407 | -107.1692
19.3444 | -107.1711
19.3481 | -107.1674
19.3519 | -107.1619
19.3556 | -107.1582
19.3593 | -107.1557
19.363 | -107.1534
19.3667 | -107.1511
19.3704 | -107.1468
19.3741 | -107.143
19.3778 | -107.1391
19.3815 | -107.1337
19.3852 | -107.1271
19.3889 | -107.122
19.3926 | -107.1156
19.3963 | -107.1094
19.4 | -107.1043
19.4037 | -107.0989
19.4074 | -107.0911
19.4111 | -107.0819
19.4148 | -107.0751
19.4185 | -107.0676
19.4222 | -107.0611
19.4259 | -107.0549
19.4296 | -107.0466
19.4333 | -107.0404
19.437 | -107.035
19.4407 | -107.0311
19.4444 | -107.0282
19.4481 | -107.0246
19.4519 | -107.0187
19.4556 | -107.0131
19.4593 | -107.0083
19.463 | -107.0039
19.4667 | -106.9992
19.4704 | -106.9952
19.4741 | -106.9918
19.4778 | -106.9905
19.4815 | -106.9876
19.4852 | -106.9851
19.4889 | -106.9852
19.4926 | -106.9839
19.4963 | -106.9838
19.5 | -106.9862
19.5037 | -106.9871
19.5074 | -106.9897
19.5111 | -106.9914
19.5148 | -106.9936
19.5185 | -106.9947
19.5222 | -106.9964
19.5259 | -107.0001
19.5296 | -107.0045
19.5333 | -107.0079
19.537 | -107.0116
19.5407 | -107.0152
19.5444 | -107.0174
19.5481 | -107.0179
19.5519 | -107.0212
19.5556 | -107.0222
19.5593 | -107.023
19.563 | -107.026
19.5667 | -107.0278
19.5704 | -107.0289
19.5741 | -107.0287
19.5778 | -107.029
19.5815 | -107.0295
19.5852 | -107.0303
19.5889 | -107.0299
19.5926 | -107.0274
19.5963 | -107.0264
19.6 | -107.0266
19.6037 | -107.0257
19.6074 | -107.0224
19.6111 | -107.022
19.6148 | -107.0225
19.6185 | -107.0179
19.6222 | -107.0129
19.6259 | -107.0132
19.6296 | -107.0141
19.6333 | -107.0094
19.637 | -107.0051
19.6407 | -107.0002
19.6444 | -106.996
19.6481 | -106.9955
19.6519 | -106.9944
19.6556 | -106.9934
19.6593 | -106.9922
19.663 | -106.9918
19.6667 | -106.9899
19.6704 | -106.9882
19.6741 | -106.9871
19.6778 | -106.9861
19.6815 | -106.9864
19.6852 | -106.9883
19.6889 | -106.989
19.6926 | -106.9887
19.6963 | -106.9888
19.7 | -106.9883
19.7037 | -106.9886
19.7074 | -106.9882
19.7111 | -106.9881
19.7148 | -106.9904
19.7185 | -106.9923
19.7222 | -106.9924
19.7259 | -106.9958
19.7296 | -106.9999
19.7333 | -107.0027
19.737 | -107.0043
19.7407 | -107.0062
19.7444 | -107.0117
19.7481 | -107.0179
19.7519 | -107.0216
19.7556 | -107.0225
19.7593 | -107.0254
19.763 | -107.0263
19.7667 | -107.0264
19.7704 | -107.0294
19.7741 | -107.0318
19.7778 | -107.0315
19.7815 | -107.0297
19.7852 | -107.0282
19.7889 | -107.0274
19.7926 | -107.0286
19.7963 | -107.031
19.8 | -107.0302
19.8037 | -107.0264
19.8074 | -107.0223
19.8111 | -107.0182
19.8148 | -107.0146
19.8185 | -107.0111
19.8222 | -107.0094
19.8259 | -107.0065
19.8296 | -107.0036
19.8333 | -106.9999
19.837 | -106.9972
19.8407 | -106.9924
19.8444 | -106.9863
19.8481 | -106.9802
19.8519 | -106.974
19.8556 | -106.9699
19.8593 | -106.9649
19.863 | -106.9608
19.8667 | -106.9546
19.8704 | -106.9461
19.8741 | -106.9392
19.8778 | -106.932
19.8815 | -106.9225
19.8852 | -106.9155
19.8889 | -106.9104
19.8926 | -106.9032
19.8963 | -106.8946
19.9 | -106.888
19.9037 | -106.8818
19.9074 | -106.8751
19.9111 | -106.8693
19.9148 | -106.8646
19.9185 | -106.8586
19.9222 | -106.8509
19.9259 | -106.8441
19.9296 | -106.8362
19.9333 | -106.8275
19.937 | -106.8216
19.9407 | -106.8167
19.9444 | -106.8124
19.9481 | -106.8082
19.9519 | -106.8041
19.9556 | -106.8002
19.9593 | -106.796
19.963 | -106.7913
19.9667 | -106.7879
19.9704 | -106.7878
19.9741 | -106.7878
19.9778 | -106.787
19.9815 | -106.7854
19.9852 | -106.7837
19.9889 | -106.7848
19.9926 | -106.7857
19.9963 | -106.7858
20 | -106.788
20.0037 | -106.7917
20.0074 | -106.7931
20.0111 | -106.7955
20.0148 | -106.8003
20.0185 | -106.806
20.0222 | -106.8108
20.0259 | -106.8136
20.0296 | -106.8174
20.0333 | -106.8229
20.037 | -106.8324
20.0407 | -106.8401
20.0444 | -106.8493
20.0481 | -106.8577
20.0519 | -106.8644
20.0556 | -106.8713
20.0593 | -106.8804
20.063 | -106.8878
20.0667 | -106.8962
20.0704 | -106.9052
20.0741 | -106.9128
20.0778 | -106.9198
20.0815 | -106.9278
20.0852 | -106.938
20.0889 | -106.9469
20.0926 | -106.9544
20.0963 | -106.96
20.1 | -106.965
20.1037 | -106.9691
20.1074 | -106.9732
20.1111 | -106.9774
20.1148 | -106.9796
20.1185 | -106.9794
20.1222 | -106.9782
20.1259 | -106.9787
20.1296 | -106.9787
20.1333 | -106.978
20.137 | -106.976
20.1407 | -106.9726
20.1444 | -106.9699
20.1481 | -106.9642
20.1519 | -106.9596
20.1556 | -106.9553
20.1593 | -106.9493
20.163 | -106.9433
20.1667 | -106.9381
20.1704 | -106.9314
20.1741 | -106.9232
20.1778 | -106.9152
20.1815 | -106.9075
20.1852 | -106.8996
20.1889 | -106.8917
20.1926 | -106.8815
20.1963 | -106.871
20.2 | -106.8633
20.2037 | -106.8562
20.2074 | -106.8507
20.2111 | -106.8435
20.2148 | -106.8373
20.2185 | -106.8329
20.2222 | -106.827
20.2259 | -106.8211
20.2296 | -106.8154
20.2333 | -106.8122
20.237 | -106.8089
20.2407 | -106.8056
20.2444 | -106.803
20.2481 | -106.8013
20.2519 | -106.7999
20.2556 | -106.7992
20.2593 | -106.7995
20.263 | -106.7993
20.2667 | -106.8002
20.2704 | -106.8012
20.2741 | -106.8031
20.2778 | -106.8047
20.2815 | -106.8093
20.2852 | -106.8135
20.2889 | -106.8185
20.2926 | -106.8246
20.2963 | -106.8306
20.3 | -106.8374
20.3037 | -106.8503
20.3074 | -106.8593
20.3111 | -106.8677
20.3148 | -106.8763
20.3185 | -106.8853
20.3222 | -106.8945
20.3259 | -106.9027
20.3296 | -106.9114
20.3333 | -106.9196
20.337 | -106.9283
20.3407 | -106.9389
20.3444 | -106.9492
20.3481 | -106.9597
20.3519 | -106.972
20.3556 | -106.9832
20.3593 | -106.9929
20.363 | -107.0023
20.3667 | -107.0102
20.3704 | -107.0184
20.3741 | -107.0266
20.3778 | -107.0355
20.3815 | -107.0438
20.3852 | -107.0522
20.3889 | -107.0592
20.3926 | -107.0643
20.3963 | -107.068
20.4 | -107.0722
20.4037 | -107.0763
20.4074 | -107.081
20.4111 | -107.084
20.4148 | -107.0854
20.4185 | -107.0859
20.4222 | -107.0862
20.4259 | -107.0853
20.4296 | -107.0836
20.4333 | -107.0792
20.437 | -107.0752
20.4407 | -107.071
20.4444 | -107.0654
20.4481 | -107.06
20.4519 | -107.0527
20.4556 | -107.0439
20.4593 | -107.0339
20.463 | -107.0246
20.4667 | -107.016
20.4704 | -107.0069
20.4741 | -106.9962
20.4778 | -106.9845
20.4815 | -106.9708
20.4852 | -106.958
20.4889 | -106.9464
20.4926 | -106.9336
20.4963 | -106.918
20.5 | -106.9051
20.5037 | -106.8923
20.5074 | -106.8785
20.5111 | -106.8659
20.5148 | -106.855
20.5185 | -106.8428
20.5222 | -106.8284
20.5259 | -106.816
20.5296 | -106.8068
20.5333 | -106.7986
20.537 | -106.7916
20.5407 | -106.7851
20.5444 | -106.7796
20.5481 | -106.7744
20.5519 | -106.771
20.5556 | -106.7687
20.5593 | -106.7677
20.563 | -106.7661
20.5667 | -106.7662
20.5704 | -106.7657
20.5741 | -106.7682
20.5778 | -106.7698
20.5815 | -106.772
20.5852 | -106.7755
20.5889 | -106.781
20.5926 | -106.7867
20.5963 | -106.7915
20.6 | -106.7979
20.6037 | -106.8085
20.6074 | -106.8198
20.6111 | -106.8303
20.6148 | -106.8407
20.6185 | -106.8502
20.6222 | -106.8627
20.6259 | -106.875
20.6296 | -106.8862
20.6333 | -106.8952
20.637 | -106.905
20.6407 | -106.9163
20.6444 | -106.9253
20.6481 | -106.936
20.6519 | -106.9454
20.6556 | -106.9544
20.6593 | -106.965
20.663 | -106.9724
20.6667 | -106.978
20.6704 | -106.9832
20.6741 | -106.9878
20.6778 | -106.9914
20.6815 | -106.9964
20.6852 | -107.0006
20.6889 | -107.001
20.6926 | -107.0007
20.6963 | -107.0001
20.7 | -106.9981
20.7037 | -106.9956
20.7074 | -106.9911
20.7111 | -106.9843
20.7148 | -106.9773
20.7185 | -106.9692
20.7222 | -106.9613
20.7259 | -106.9544
20.7296 | -106.9445
20.7333 | -106.9338
20.737 | -106.9251
20.7407 | -106.9143
20.7444 | -106.9018
20.7481 | -106.89
20.7519 | -106.8788
20.7556 | -106.8668
20.7593 | -106.8519
20.763 | -106.8374
20.7667 | -106.8249
20.7704 | -106.8118
20.7741 | -106.7978
20.7778 | -106.7842
20.7815 | -106.7701
20.7852 | -106.7564
20.7889 | -106.7456
20.7926 | -106.7332
20.7963 | -106.7196
20.8 | -106.7079
20.8037 | -106.6962
20.8074 | -106.6819
20.8111 | -106.6673
20.8148 | -106.6558
20.8185 | -106.643
20.8222 | -106.6299
20.8259 | -106.616
20.8296 | -106.6022
20.8333 | -106.5885
20.837 | -106.5772
20.8407 | -106.5641
20.8444 | -106.5534
20.8481 | -106.5428
20.8519 | -106.5338
20.8556 | -106.5261
20.8593 | -106.5194
20.863 | -106.5126
20.8667 | -106.5066
20.8704 | -106.4988
20.8741 | -106.4941
20.8778 | -106.4915
20.8815 | -106.4884
20.8852 | -106.4839
20.8889 | -106.4817
20.8926 | -106.4824
20.8963 | -106.4824
20.9 | -106.4811
20.9037 | -106.4827
20.9074 | -106.4861
20.9111 | -106.491
20.9148 | -106.4979
20.9185 | -106.5016
20.9222 | -106.506
20.9259 | -106.5124
20.9296 | -106.52
20.9333 | -106.5281
20.937 | -106.5363
20.9407 | -106.5449
20.9444 | -106.5524
20.9481 | -106.5623
20.9519 | -106.574
20.9556 | -106.5843
20.9593 | -106.5948
20.963 | -106.6071
20.9667 | -106.6193
20.9704 | -106.6321
20.9741 | -106.646
20.9778 | -106.6576
20.9815 | -106.6688
20.9852 | -106.6817
20.9889 | -106.693
20.9926 | -106.704
20.9963 | -106.7152
21 | -106.7275
21.0037 | -106.7388
21.0074 | -106.7506
21.0111 | -106.7591
21.0148 | -106.7682
21.0185 | -106.7758
21.0222 | -106.7839
21.0259 | -106.7942
21.0296 | -106.8035
21.0333 | -106.8103
21.037 | -106.8166
21.0407 | -106.8239
21.0444 | -106.8315
21.0481 | -106.8397
21.0519 | -106.8448
21.0556 | -106.8494
21.0593 | -106.8535
21.063 | -106.8545
21.0667 | -106.8561
21.0704 | -106.8597
21.0741 | -106.8628
21.0778 | -106.8646
21.0815 | -106.8656
21.0852 | -106.8645
21.0889 | -106.8642
21.0926 | -106.8624
21.0963 | -106.8579
21.1 | -106.854
21.1037 | -106.8529
21.1074 | -106.852
21.1111 | -106.8487
21.1148 | -106.8452
21.1185 | -106.8388
21.1222 | -106.8319
21.1259 | -106.8256
21.1296 | -106.8198
21.1333 | -106.8144
21.137 | -106.8097
21.1407 | -106.8035
21.1444 | -106.7946
21.1481 | -106.7849
21.1519 | -106.7752
21.1556 | -106.7657
21.1593 | -106.7556
21.163 | -106.7477
21.1667 | -106.738
21.1704 | -106.7309
21.1741 | -106.7248
21.1778 | -106.7186
21.1815 | -106.7136
21.1852 | -106.7079
21.1889 | -106.7004
21.1926 | -106.6934
21.1963 | -106.6897
21.2 | -106.6856
21.2037 | -106.6795
21.2074 | -106.6746
21.2111 | -106.6719
21.2148 | -106.6688
21.2185 | -106.6689
21.2222 | -106.6713
21.2259 | -106.673
21.2296 | -106.6718
21.2333 | -106.6706
21.237 | -106.6705
21.2407 | -106.6735
21.2444 | -106.6749
21.2481 | -106.6752
21.2519 | -106.6764
21.2556 | -106.6797
21.2593 | -106.6828
21.263 | -106.6859
21.2667 | -106.6878
21.2704 | -106.6882
21.2741 | -106.6881
21.2778 | -106.691
21.2815 | -106.6946
21.2852 | -106.6964
21.2889 | -106.6992
21.2926 | -106.7017
21.2963 | -106.7034
21.3 | -106.7046
21.3037 | -106.7076
21.3074 | -106.7081
21.3111 | -106.7098
21.3148 | -106.7109
21.3185 | -106.7135
21.3222 | -106.7151
21.3259 | -106.7137
21.3296 | -106.7099
21.3333 | -106.7089
21.337 | -106.7096
21.3407 | -106.709
21.3444 | -106.708
21.3481 | -106.7054
21.3519 | -106.7038
21.3556 | -106.7029
21.3593 | -106.7019
21.363 | -106.7031
21.3667 | -106.7023
21.3704 | -106.6982
21.3741 | -106.6972
21.3778 | -106.6943
21.3815 | -106.6911
21.3852 | -106.6903
21.3889 | -106.6916
21.3926 | -106.692
21.3963 | -106.6932
21.4 | -106.6961
21.4037 | -106.6984
21.4074 | -106.6996
21.4111 | -106.7018
21.4148 | -106.7048
21.4185 | -106.7075
21.4222 | -106.7105
21.4259 | -106.7134
21.4296 | -106.7157
21.4333 | -106.7192
21.437 | -106.722
21.4407 | -106.7239
21.4444 | -106.7279
21.4481 | -106.7338
21.4519 | -106.7416
21.4556 | -106.749
21.4593 | -106.7554
21.463 | -106.7614
21.4667 | -106.7671
21.4704 | -106.775
21.4741 | -106.7817
21.4778 | -106.787
21.4815 | -106.7922
21.4852 | -106.7993
21.4889 | -106.8072
21.4926 | -106.8151
21.4963 | -106.8217
21.5 | -106.8267
21.5037 | -106.832
21.5074 | -106.8385
21.5111 | -106.8443
21.5148 | -106.8474
21.5185 | -106.8533
21.5222 | -106.8577
21.5259 | -106.8594
21.5296 | -106.8617
21.5333 | -106.866
21.537 | -106.8702
21.5407 | -106.8731
21.5444 | -106.8746
21.5481 | -106.877
21.5519 | -106.8815
21.5556 | -106.8838
21.5593 | -106.8836
21.563 | -106.8845
21.5667 | -106.8859
21.5704 | -106.8875
21.5741 | -106.8878
21.5778 | -106.8894
21.5815 | -106.8892
21.5852 | -106.8882
21.5889 | -106.888
21.5926 | -106.8847
21.5963 | -106.8798
21.6 | -106.8761
21.6037 | -106.8728
21.6074 | -106.8686
21.6111 | -106.865
21.6148 | -106.861
21.6185 | -106.8566
21.6222 | -106.8518
21.6259 | -106.8463
21.6296 | -106.8408
21.6333 | -106.8395
21.637 | -106.8366
21.6407 | -106.833
21.6444 | -106.8312
21.6481 | -106.8281
21.6519 | -106.8239
21.6556 | -106.8215
21.6593 | -106.8192
21.663 | -106.8171
21.6667 | -106.8118
21.6704 | -106.8084
21.6741 | -106.8071
21.6778 | -106.8038
21.6815 | -106.7979
21.6852 | -106.7921
21.6889 | -106.7869
21.6926 | -106.781
21.6963 | -106.7747
21.7 | -106.7673
21.7037 | -106.7613
21.7074 | -106.7558
21.7111 | -106.7498
21.7148 | -106.7433
21.7185 | -106.7361
21.7222 | -106.7285
21.7259 | -106.7212
21.7296 | -106.7145
21.7333 | -106.7066
21.737 | -106.6983
21.7407 | -106.6902
21.7444 | -106.6813
21.7481 | -106.674
21.7519 | -106.6649
21.7556 | -106.6565
21.7593 | -106.648
21.763 | -106.6388
21.7667 | -106.6287
21.7704 | -106.6199
21.7741 | -106.611
21.7778 | -106.6041
21.7815 | -106.5978
21.7852 | -106.5902
21.7889 | -106.5829
21.7926 | -106.5751
21.7963 | -106.5678
21.8 | -106.5606
21.8037 | -106.554
21.8074 | -106.549
21.8111 | -106.5448
21.8148 | -106.5383
21.8185 | -106.5331
21.8222 | -106.5297
21.8259 | -106.5248
21.8296 | -106.5199
21.8333 | -106.5181
21.837 | -106.5161
21.8407 | -106.5134
21.8444 | -106.5089
21.8481 | -106.5049
21.8519 | -106.5032
21.8556 | -106.5015
21.8593 | -106.4968
21.863 | -106.4929
21.8667 | -106.4913
21.8704 | -106.4895
21.8741 | -106.4863
21.8778 | -106.4827
21.8815 | -106.4813
21.8852 | -106.4787
21.8889 | -106.476
21.8926 | -106.4736
21.8963 | -106.4727
21.9 | -106.4721
21.9037 | -106.4715
21.9074 | -106.4675
21.9111 | -106.4655
21.9148 | -106.4635
21.9185 | -106.462
21.9222 | -106.4586
21.9259 | -106.4561
21.9296 | -106.4522
21.9333 | -106.4484
21.937 | -106.4441
21.9407 | -106.4406
21.9444 | -106.4354
21.9481 | -106.4323
21.9519 | -106.4296
21.9556 | -106.4254
21.9593 | -106.4233
21.963 | -106.4235
21.9667 | -106.4228
21.9704 | -106.4221
21.9741 | -106.4205
21.9778 | -106.4208
21.9815 | -106.4198
21.9852 | -106.4198
21.9889 | -106.4201
21.9926 | -106.4201
21.9963 | -106.4199
22 | -106.4198
22.0037 | -106.4193
22.0074 | -106.4195
22.0111 | -106.421
22.0148 | -106.4212
22.0185 | -106.4199
22.0222 | -106.4196
22.0259 | -106.4202
22.0296 | -106.4223
22.0333 | -106.4231
22.037 | -106.4242
22.0407 | -106.4255
22.0444 | -106.426
22.0481 | -106.4258
22.0519 | -106.4268
22.0556 | -106.4271
22.0593 | -106.4267
22.063 | -106.4283
22.0667 | -106.4284
22.0704 | -106.4293
22.0741 | -106.4303
22.0778 | -106.4323
22.0815 | -106.4352
22.0852 | -106.4368
22.0889 | -106.4363
22.0926 | -106.4343
22.0963 | -106.4355
22.1 | -106.4366
22.1037 | -106.4378
22.1074 | -106.4391
22.1111 | -106.4425
22.1148 | -106.4462
22.1185 | -106.4494
22.1222 | -106.4519
22.1259 | -106.4563
22.1296 | -106.4598
22.1333 | -106.4624
22.137 | -106.4638
22.1407 | -106.4673
22.1444 | -106.4711
22.1481 | -106.4754
22.1519 | -106.4808
22.1556 | -106.4872
22.1593 | -106.4918
22.163 | -106.4937
22.1667 | -106.4958
22.1704 | -106.4998
22.1741 | -106.5039
22.1778 | -106.5068
22.1815 | -106.5088
22.1852 | -106.5124
22.1889 | -106.5183
22.1926 | -106.522
22.1963 | -106.5235
22.2 | -106.5269
22.2037 | -106.5319
22.2074 | -106.5359
22.2111 | -106.5402
22.2148 | -106.5429
22.2185 | -106.5446
22.2222 | -106.5489
22.2259 | -106.5532
22.2296 | -106.5543
22.2333 | -106.5553
22.237 | -106.5578
22.2407 | -106.557
22.2444 | -106.5569
22.2481 | -106.5571
22.2519 | -106.559
22.2556 | -106.56
22.2593 | -106.5596
22.263 | -106.5563
22.2667 | -106.5544
22.2704 | -106.5524
22.2741 | -106.5482
22.2778 | -106.5443
22.2815 | -106.5416
22.2852 | -106.5404
22.2889 | -106.5367
22.2926 | -106.5346
22.2963 | -106.5342
22.3 | -106.5321
22.3037 | -106.5305
22.3074 | -106.5299
22.3111 | -106.5282
22.3148 | -106.5245
22.3185 | -106.5204
22.3222 | -106.5166
22.3259 | -106.514
22.3296 | -106.5125
22.3333 | -106.5112
22.337 | -106.508
22.3407 | -106.5012
22.3444 | -106.4941
22.3481 | -106.4875
22.3519 | -106.479
22.3556 | -106.4739
22.3593 | -106.4704
22.363 | -106.4634
22.3667 | -106.4543
22.3704 | -106.4492
22.3741 | -106.4425
22.3778 | -106.4333
22.3815 | -106.4245
22.3852 | -106.4141
22.3889 | -106.4025
22.3926 | -106.3917
22.3963 | -106.3806
22.4 | -106.3688
22.4037 | -106.3557
22.4074 | -106.3431
22.4111 | -106.3287
22.4148 | -106.3138
22.4185 | -106.297
22.4222 | -106.2787
22.4259 | -106.2587
22.4296 | -106.2377
22.4333 | -106.214
22.437 | -106.1892
22.4407 | -106.1572
22.4444 | -106.133
22.4481 | -106.1091
22.4519 | -106.0849
22.4556 | -106.061
22.4593 | -106.0376
22.463 | -106.0138
22.4667 | -105.9895
22.4704 | -105.9644
22.4741 | -105.9398
22.4778 | -105.9178
22.4815 | -105.8979
22.4852 | -105.8784
22.4889 | -105.8587
22.4926 | -105.8397
22.4963 | -105.8219
22.5 | -105.8057
22.5037 | -105.7907
22.5074 | -105.7763
22.5111 | -105.7634
22.5148 | -105.7522
22.5185 | -105.7419
22.5222 | -105.7326
22.5259 | -105.7248
22.5296 | -105.7207
22.5333 | -105.7203
22.537 | -105.7213
22.5407 | -105.7218
22.5444 | -105.7252
22.5481 | -105.7309
22.5519 | -105.7387
22.5556 | -105.7474
22.5593 | -105.7574
22.563 | -105.7694
22.5667 | -105.7827
22.5704 | -105.7982
22.5741 | -105.8166
22.5778 | -105.8364
22.5815 | -105.8573
22.5852 | -105.8791
22.5889 | -105.9032
22.5926 | -105.9289
22.5963 | -105.9556
22.6 | -105.9829
22.6037 | -106.0091
22.6074 | -106.0372
22.6111 | -106.0654
22.6148 | -106.0926
22.6185 | -106.1219
22.6222 | -106.1516
22.6259 | -106.181
22.6296 | -106.2102
22.6333 | -106.2419
22.637 | -106.2744
22.6407 | -106.3057
22.6444 | -106.336
22.6481 | -106.3655
22.6519 | -106.3949
22.6556 | -106.4233
22.6593 | -106.4528
22.663 | -106.4815
22.6667 | -106.5098
22.6704 | -106.5375
22.6741 | -106.5597
22.6778 | -106.5788
22.6815 | -106.5992
22.6852 | -106.6184
22.6889 | -106.6352
22.6926 | -106.6509
22.6963 | -106.665
22.7 | -106.6768
22.7037 | -106.688
22.7074 | -106.7003
22.7111 | -106.7094
22.7148 | -106.7162
22.7185 | -106.7224
22.7222 | -106.7286
22.7259 | -106.7344
22.7296 | -106.739
22.7333 | -106.7394
22.737 | -106.7395
22.7407 | -106.7393
22.7444 | -106.7392
22.7481 | -106.7398
22.7519 | -106.7423
22.7556 | -106.7425
22.7593 | -106.7427
22.763 | -106.7409
22.7667 | -106.739
22.7704 | -106.7361
22.7741 | -106.7327
22.7778 | -106.7289
22.7815 | -106.7237
22.7852 | -106.7181
22.7889 | -106.7122
22.7926 | -106.7077
22.7963 | -106.7021
22.8 | -106.6932
22.8037 | -106.6849
22.8074 | -106.6794
22.8111 | -106.6748
22.8148 | -106.6695
22.8185 | -106.6644
22.8222 | -106.66
22.8259 | -106.6557
22.8296 | -106.6505
22.8333 | -106.6459
22.837 | -106.6422
22.8407 | -106.6396
22.8444 | -106.6372
22.8481 | -106.6339
22.8519 | -106.6302
22.8556 | -106.6268
22.8593 | -106.6242
22.863 | -106.6208
22.8667 | -106.6162
22.8704 | -106.6152
22.8741 | -106.6138
22.8778 | -106.6113
22.8815 | -106.6122
22.8852 | -106.6135
22.8889 | -106.6121
22.8926 | -106.6101
22.8963 | -106.6085
22.9 | -106.6091
22.9037 | -106.6101
22.9074 | -106.6114
22.9111 | -106.6127
22.9148 | -106.6144
22.9185 | -106.6144
22.9222 | -106.6152
22.9259 | -106.617
22.9296 | -106.62
22.9333 | -106.6212
22.937 | -106.6234
22.9407 | -106.6247
22.9444 | -106.6277
22.9481 | -106.6306
22.9519 | -106.6344
22.9556 | -106.6355
22.9593 | -106.6364
22.963 | -106.6374
22.9667 | -106.6366
22.9704 | -106.6366
22.9741 | -106.6369
22.9778 | -106.6379
22.9815 | -106.639
22.9852 | -106.6413
22.9889 | -106.6418
22.9926 | -106.6402
22.9963 | -106.639
23 | -106.6404
23.0037 | -106.643
23.0074 | -106.645
23.0111 | -106.6441
23.0148 | -106.6418
23.0185 | -106.6404
23.0222 | -106.6395
23.0259 | -106.6403
23.0296 | -106.6402
23.0333 | -106.6383
23.037 | -106.6363
23.0407 | -106.6325
23.0444 | -106.6295
23.0481 | -106.6268
23.0519 | -106.6244
23.0556 | -106.6227
23.0593 | -106.6232
23.063 | -106.6234
23.0667 | -106.6227
23.0704 | -106.6221
23.0741 | -106.6194
23.0778 | -106.6158
23.0815 | -106.6162
23.0852 | -106.6142
23.0889 | -106.6108
23.0926 | -106.6081
23.0963 | -106.6054
23.1 | -106.6041
23.1037 | -106.603
23.1074 | -106.6025
23.1111 | -106.6014
23.1148 | -106.6007
23.1185 | -106.6006
23.1222 | -106.6
23.1259 | -106.6003
23.1296 | -106.5991
23.1333 | -106.5964
23.137 | -106.5939
23.1407 | -106.5917
23.1444 | -106.5896
23.1481 | -106.5863
23.1519 | -106.5812
23.1556 | -106.5777
23.1593 | -106.5748
23.163 | -106.5729
23.1667 | -106.5701
23.1704 | -106.568
23.1741 | -106.5675
23.1778 | -106.5656
23.1815 | -106.5633
23.1852 | -106.5615
23.1889 | -106.5597
23.1926 | -106.5576
23.1963 | -106.5572
23.2 | -106.5568
23.2037 | -106.5574
23.2074 | -106.5571
23.2111 | -106.5545
23.2148 | -106.5525
23.2185 | -106.5509
23.2222 | -106.5497
23.2259 | -106.5483
23.2296 | -106.5483
23.2333 | -106.548
23.237 | -106.5477
23.2407 | -106.5478
23.2444 | -106.548
23.2481 | -106.5478
23.2519 | -106.5481
23.2556 | -106.5488
23.2593 | -106.5484
23.263 | -106.5469
23.2667 | -106.5462
23.2704 | -106.5445
23.2741 | -106.5431
23.2778 | -106.5419
23.2815 | -106.5409
23.2852 | -106.5374
23.2889 | -106.5344
23.2926 | -106.5327
23.2963 | -106.5297
23.3 | -106.5256
23.3037 | -106.5233
23.3074 | -106.5214
23.3111 | -106.5203
23.3148 | -106.5163
23.3185 | -106.5118
23.3222 | -106.5085
23.3259 | -106.5063
23.3296 | -106.4994
23.3333 | -106.4939
23.337 | -106.4891
23.3407 | -106.485
23.3444 | -106.481
23.3481 | -106.4778
23.3519 | -106.4743
23.3556 | -106.4695
23.3593 | -106.4654
23.363 | -106.461
23.3667 | -106.457
23.3704 | -106.4543
23.3741 | -106.452
23.3778 | -106.4489
23.3815 | -106.444
23.3852 | -106.4396
23.3889 | -106.4364
23.3926 | -106.4349
23.3963 | -106.4342
23.4 | -106.4321
23.4037 | -106.4304
23.4074 | -106.4298
23.4111 | -106.429
23.4148 | -106.4304
23.4185 | -106.4318
23.4222 | -106.4331
23.4259 | -106.4347
23.4296 | -106.4363
23.4333 | -106.4381
23.437 | -106.44
23.4407 | -106.4434
23.4444 | -106.4474
23.4481 | -106.4515
23.4519 | -106.4568
23.4556 | -106.4629
23.4593 | -106.4698
23.463 | -106.4772
23.4667 | -106.4845
23.4704 | -106.4905
23.4741 | -106.4969
23.4778 | -106.5068
23.4815 | -106.5178
23.4852 | -106.5279
23.4889 | -106.5383
23.4926 | -106.5466
23.4963 | -106.5549
23.5 | -106.5645
23.5037 | -106.5733
23.5074 | -106.5812
23.5111 | -106.589
23.5148 | -106.5971
23.5185 | -106.6057
23.5222 | -106.6152
23.5259 | -106.624
23.5296 | -106.6305
23.5333 | -106.6358
23.537 | -106.643
23.5407 | -106.65
23.5444 | -106.6569
23.5481 | -106.6635
23.5519 | -106.6684
23.5556 | -106.6733
23.5593 | -106.6775
23.563 | -106.6809
23.5667 | -106.6835
23.5704 | -106.6882
23.5741 | -106.6924
23.5778 | -106.6963
23.5815 | -106.7009
23.5852 | -106.7047
23.5889 | -106.7073
23.5926 | -106.712
23.5963 | -106.7166
23.6 | -106.7199
23.6037 | -106.7222
23.6074 | -106.7257
23.6111 | -106.7287
23.6148 | -106.7314
23.6185 | -106.7353
23.6222 | -106.7377
23.6259 | -106.7402
23.6296 | -106.7445
23.6333 | -106.749
23.637 | -106.7497
23.6407 | -106.7497
23.6444 | -106.7505
23.6481 | -106.7551
23.6519 | -106.7605
23.6556 | -106.7657
23.6593 | -106.7689
23.663 | -106.774
23.6667 | -106.7774
23.6704 | -106.781
23.6741 | -106.7854
23.6778 | -106.7878
23.6815 | -106.7896
23.6852 | -106.7918
23.6889 | -106.7935
23.6926 | -106.796
23.6963 | -106.7978
23.7 | -106.7989
23.7037 | -106.8015
23.7074 | -106.8019
23.7111 | -106.7999
23.7148 | -106.801
23.7185 | -106.8009
23.7222 | -106.7996
23.7259 | -106.7981
23.7296 | -106.7963
23.7333 | -106.7948
23.737 | -106.7946
23.7407 | -106.7934
23.7444 | -106.7919
23.7481 | -106.79
23.7519 | -106.7889
23.7556 | -106.7872
23.7593 | -106.785
23.763 | -106.7827
23.7667 | -106.7806
23.7704 | -106.7792
23.7741 | -106.7771
23.7778 | -106.7742
23.7815 | -106.7705
23.7852 | -106.7682
23.7889 | -106.7655
23.7926 | -106.7623
23.7963 | -106.7585
23.8 | -106.7558
23.8037 | -106.7557
23.8074 | -106.7553
23.8111 | -106.7519
23.8148 | -106.7477
23.8185 | -106.7454
23.8222 | -106.7439
23.8259 | -106.7435
23.8296 | -106.7415
23.8333 | -106.7374
23.837 | -106.7361
23.8407 | -106.7356
23.8444 | -106.7344
23.8481 | -106.7331
23.8519 | -106.7322
23.8556 | -106.7323
23.8593 | -106.7324
23.863 | -106.7331
23.8667 | -106.7333
23.8704 | -106.7346
23.8741 | -106.7346
23.8778 | -106.734
23.8815 | -106.7345
23.8852 | -106.7353
23.8889 | -106.7367
23.8926 | -106.7379
23.8963 | -106.7389
23.9 | -106.7401
23.9037 | -106.7411
23.9074 | -106.7402
23.9111 | -106.741
23.9148 | -106.744
23.9185 | -106.7464
23.9222 | -106.7473
23.9259 | -106.7492
23.9296 | -106.7523
23.9333 | -106.7556
23.937 | -106.7582
23.9407 | -106.7617
23.9444 | -106.7634
23.9481 | -106.7661
23.9519 | -106.7711
23.9556 | -106.7757
23.9593 | -106.7811
23.963 | -106.7863
23.9667 | -106.789
23.9704 | -106.7922
23.9741 | -106.7963
23.9778 | -106.7995
23.9815 | -106.8036
23.9852 | -106.8057
23.9889 | -106.8085
23.9926 | -106.8114
23.9963 | -106.8152
24 | -106.8193
24.0037 | -106.8221
24.0074 | -106.8236
24.0111 | -106.8253
24.0148 | -106.8282
24.0185 | -106.8302
24.0222 | -106.8334
24.0259 | -106.8358
24.0296 | -106.8359
24.0333 | -106.8388
24.037 | -106.8424
24.0407 | -106.8444
24.0444 | -106.8449
24.0481 | -106.8473
24.0519 | -106.8501
24.0556 | -106.8505
24.0593 | -106.8501
24.063 | -106.8511
24.0667 | -106.8494
24.0704 | -106.8495
24.0741 | -106.8492
24.0778 | -106.848
24.0815 | -106.8473
24.0852 | -106.8467
24.0889 | -106.8457
24.0926 | -106.8441
24.0963 | -106.8424
24.1 | -106.8409
24.1037 | -106.84
24.1074 | -106.839
24.1111 | -106.8373
24.1148 | -106.8354
24.1185 | -106.8348
24.1222 | -106.8351
24.1259 | -106.8337
24.1296 | -106.8309
24.1333 | -106.8291
24.137 | -106.8294
24.1407 | -106.8283
24.1444 | -106.8256
24.1481 | -106.8228
24.1519 | -106.8213
24.1556 | -106.8193
24.1593 | -106.8157
24.163 | -106.8104
24.1667 | -106.805
24.1704 | -106.8034
24.1741 | -106.8019
24.1778 | -106.7988
24.1815 | -106.7951
24.1852 | -106.7916
24.1889 | -106.7884
24.1926 | -106.7833
24.1963 | -106.7778
24.2 | -106.7714
24.2037 | -106.7645
24.2074 | -106.7586
24.2111 | -106.7532
24.2148 | -106.7479
24.2185 | -106.7423
24.2222 | -106.736
24.2259 | -106.7299
24.2296 | -106.7242
24.2333 | -106.7182
24.237 | -106.7121
24.2407 | -106.7065
24.2444 | -106.6999
24.2481 | -106.6902
24.2519 | -106.679
24.2556 | -106.6663
24.2593 | -106.6283
24.263 | -106.5563
24.2667 | -106.4663
24.2704 | -106.3966
24.2741 | -106.3269
24.2778 | -106.2506
24.2815 | -106.1654
24.2852 | -106.0792
24.2889 | -105.9902
24.2926 | -105.9072
24.2963 | -105.83
24.3 | -105.7488
24.3037 | -105.6618
24.3074 | -105.572
24.3111 | -105.4793
24.3148 | -105.3828
24.3185 | -105.2829
24.3222 | -105.1791
24.3259 | -105.0721
24.3296 | -104.9621
24.3333 | -104.8518
24.337 | -104.7404
24.3407 | -104.6273
24.3444 | -104.5134
24.3481 | -104.3999
24.3519 | -104.2855
24.3556 | -104.1718
24.3593 | -104.0595
24.363 | -103.9475
24.3667 | -103.8376
24.3704 | -103.7307
24.3741 | -103.6258
24.3778 | -103.5239
24.3815 | -103.4255
24.3852 | -103.3313
24.3889 | -103.2415
24.3926 | -103.156
24.3963 | -103.0756
24.4 | -103.0006
24.4037 | -102.931
24.4074 | -102.8667
24.4111 | -102.8085
24.4148 | -102.7558
24.4185 | -102.7093
24.4222 | -102.6695
24.4259 | -102.6358
24.4296 | -102.6081
24.4333 | -102.5871
24.437 | -102.5731
24.4407 | -102.5659
24.4444 | -102.5651
24.4481 | -102.5708
24.4519 | -102.5838
24.4556 | -102.604
24.4593 | -102.6307
24.463 | -102.6627
24.4667 | -102.7015
24.4704 | -102.7477
24.4741 | -102.8013
24.4778 | -102.8607
24.4815 | -102.9257
24.4852 | -102.9965
24.4889 | -103.0731
24.4926 | -103.1549
24.4963 | -103.2424
24.5 | -103.3347
24.5037 | -103.4315
24.5074 | -103.5325
24.5111 | -103.6379
24.5148 | -103.7451
24.5185 | -103.8559
24.5222 | -103.9704
24.5259 | -104.0871
24.5296 | -104.206
24.5333 | -104.3258
24.537 | -104.447
24.5407 | -104.568
24.5444 | -104.6892
24.5481 | -104.8094
24.5519 | -104.9289
24.5556 | -105.046
24.5593 | -105.1579
24.563 | -105.264
24.5667 | -105.3584
24.5704 | -105.4297
24.5741 | -105.473
24.5778 | -105.3893
24.5815 | -105.3859
24.5852 | -105.3672
24.5889 | -105.3542
24.5926 | -105.3497
24.5963 | -105.3444
24.6 | -105.3298
24.6037 | -105.3029
24.6074 | -105.2801
24.6111 | -105.2624
24.6148 | -105.236
24.6185 | -105.1775
24.6222 | -105.1055
24.6259 | -105.0262
24.6296 | -104.9396
24.6333 | -104.8417
24.637 | -104.7336
24.6407 | -104.6109
24.6444 | -104.478
24.6481 | -104.3428
24.6519 | -104.2068
24.6556 | -104.0698
24.6593 | -103.9321
24.663 | -103.7952
24.6667 | -103.6586
24.6704 | -103.5245
24.6741 | -103.3932
24.6778 | -103.2638
24.6815 | -103.1367
24.6852 | -103.0138
24.6889 | -102.8954
24.6926 | -102.7818
24.6963 | -102.672
24.7 | -102.5676
24.7037 | -102.4684
24.7074 | -102.3748
24.7111 | -102.2874
24.7148 | -102.2058
24.7185 | -102.1299
24.7222 | -102.0601
24.7259 | -101.9971
24.7296 | -101.9411
24.7333 | -101.8917
24.737 | -101.8491
24.7407 | -101.8132
24.7444 | -101.7844
24.7481 | -101.7631
24.7519 | -101.7489
24.7556 | -101.742
24.7593 | -101.742
24.763 | -101.7493
24.7667 | -101.7642
24.7704 | -101.7859
24.7741 | -101.8142
24.7778 | -101.8496
24.7815 | -101.8924
24.7852 | -101.9419
24.7889 | -101.9981
24.7926 | -102.0605
24.7963 | -102.1289
24.8 | -102.2035
24.8037 | -102.2837
24.8074 | -102.3701
24.8111 | -102.4632
24.8148 | -102.5614
24.8185 | -102.6636
24.8222 | -102.7699
24.8259 | -102.8805
24.8296 | -102.995
24.8333 | -103.1135
24.837 | -103.2359
24.8407 | -103.3618
24.8444 | -103.4894
24.8481 | -103.6189
24.8519 | -103.7503
24.8556 | -103.883
24.8593 | -104.0157
24.863 | -104.1499
24.8667 | -104.283
24.8704 | -104.4152
24.8741 | -104.5468
24.8778 | -104.6764
24.8815 | -104.8049
24.8852 | -104.9309
24.8889 | -105.054
24.8926 | -105.1742
24.8963 | -105.2916
24.9 | -105.404
24.9037 | -105.5135
24.9074 | -105.6215
24.9111 | -105.7266
24.9148 | -105.8342
24.9185 | -105.9526
24.9222 | -106.0802
24.9259 | -106.2018
24.9296 | -106.3106
24.9333 | -106.4282
24.937 | -106.5361
24.9407 | -106.6097
24.9444 | -106.6495
24.9481 | -106.6719
24.9519 | -106.6889
24.9556 | -106.7031
24.9593 | -106.7178
24.963 | -106.7343
24.9667 | -106.7444
24.9704 | -106.7469
24.9741 | -106.7482
24.9778 | -106.7448
24.9815 | -106.7374
24.9852 | -106.7288
24.9889 | -106.7206
24.9926 | -106.7107
24.9963 | -106.6998
25 | -106.6874
25.0037 | -106.6751
25.0074 | -106.6621
25.0111 | -106.6485
25.0148 | -106.6359
25.0185 | -106.6245
25.0222 | -106.6085
25.0259 | -106.5899
25.0296 | -106.5732
25.0333 | -106.5559
25.037 | -106.5368
25.0407 | -106.5181
25.0444 | -106.4988
25.0481 | -106.4798
25.0519 | -106.4604
25.0556 | -106.4413
25.0593 | -106.4217
25.063 | -106.4036
25.0667 | -106.3864
25.0704 | -106.3702
25.0741 | -106.3552
25.0778 | -106.3392
25.0815 | -106.3236
25.0852 | -106.3103
25.0889 | -106.2975
25.0926 | -106.2854
25.0963 | -106.2726
25.1 | -106.2609
25.1037 | -106.2496
25.1074 | -106.2393
25.1111 | -106.2372
25.1148 | -106.2303
25.1185 | -106.2226
25.1222 | -106.2144
25.1259 | -106.2079
25.1296 | -106.2043
25.1333 | -106.203
25.137 | -106.2015
25.1407 | -106.1998
25.1444 | -106.1987
25.1481 | -106.2003
25.1519 | -106.2034
25.1556 | -106.2061
25.1593 | -106.2102
25.163 | -106.2142
25.1667 | -106.2193
25.1704 | -106.2246
25.1741 | -106.2289
25.1778 | -106.2348
25.1815 | -106.2436
25.1852 | -106.2525
25.1889 | -106.2624
25.1926 | -106.2735
25.1963 | -106.2817
25.2 | -106.2912
25.2037 | -106.3027
25.2074 | -106.3145
25.2111 | -106.327
25.2148 | -106.3398
25.2185 | -106.3523
25.2222 | -106.3666
25.2259 | -106.379
25.2296 | -106.3878
25.2333 | -106.3999
25.237 | -106.412
25.2407 | -106.422
25.2444 | -106.4325
25.2481 | -106.4439
25.2519 | -106.4543
25.2556 | -106.4647
25.2593 | -106.4732
25.263 | -106.4814
25.2667 | -106.4908
25.2704 | -106.4998
25.2741 | -106.5066
25.2778 | -106.5133
25.2815 | -106.5184
25.2852 | -106.523
25.2889 | -106.5291
25.2926 | -106.5347
25.2963 | -106.5379
25.3 | -106.5413
25.3037 | -106.5441
25.3074 | -106.5455
25.3111 | -106.5482
25.3148 | -106.5444
25.3185 | -106.538
25.3222 | -106.5365
25.3259 | -106.5365
25.3296 | -106.533
25.3333 | -106.5296
25.337 | -106.5273
25.3407 | -106.5234
25.3444 | -106.5175
25.3481 | -106.5123
25.3519 | -106.5073
25.3556 | -106.5027
25.3593 | -106.496
25.363 | -106.4886
25.3667 | -106.4816
25.3704 | -106.4713
25.3741 | -106.4582
25.3778 | -106.4483
25.3815 | -106.4412
25.3852 | -106.4337
25.3889 | -106.4261
25.3926 | -106.418
25.3963 | -106.4098
25.4 | -106.4029
25.4037 | -106.3958
25.4074 | -106.3885
25.4111 | -106.3806
25.4148 | -106.3721
25.4185 | -106.3617
25.4222 | -106.3532
25.4259 | -106.3464
25.4296 | -106.3389
25.4333 | -106.3331
25.437 | -106.3274
25.4407 | -106.3207
25.4444 | -106.3132
25.4481 | -106.3063
25.4519 | -106.3021
25.4556 | -106.2962
25.4593 | -106.2907
25.463 | -106.285
25.4667 | -106.2799
25.4704 | -106.2764
25.4741 | -106.2718
25.4778 | -106.2685
25.4815 | -106.265
25.4852 | -106.2611
25.4889 | -106.2568
25.4926 | -106.2525
25.4963 | -106.2488
25.5 | -106.2439
25.5037 | -106.2401
25.5074 | -106.2378
25.5111 | -106.2352
25.5148 | -106.2339
25.5185 | -106.2332
25.5222 | -106.2336
25.5259 | -106.234
25.5296 | -106.2347
25.5333 | -106.238
25.537 | -106.2407
25.5407 | -106.2427
25.5444 | -106.2473
25.5481 | -106.2524
25.5519 | -106.2561
25.5556 | -106.2601
25.5593 | -106.2652
25.563 | -106.2724
25.5667 | -106.279
25.5704 | -106.2853
25.5741 | -106.292
25.5778 | -106.2981
25.5815 | -106.3032
25.5852 | -106.3105
25.5889 | -106.3204
25.5926 | -106.3282
25.5963 | -106.3377
25.6 | -106.3481
25.6037 | -106.358
25.6074 | -106.3679
25.6111 | -106.3775
25.6148 | -106.3866
25.6185 | -106.3967
25.6222 | -106.4075
25.6259 | -106.4173
25.6296 | -106.4263
25.6333 | -106.437
25.637 | -106.449
25.6407 | -106.461
25.6444 | -106.4747
25.6481 | -106.4869
25.6519 | -106.499
25.6556 | -106.5097
25.6593 | -106.5211
25.663 | -106.5334
25.6667 | -106.5458
25.6704 | -106.5586
25.6741 | -106.5703
25.6778 | -106.583
25.6815 | -106.5962
25.6852 | -106.6083
25.6889 | -106.6201
25.6926 | -106.6322
25.6963 | -106.6435
25.7 | -106.6559
25.7037 | -106.6688
25.7074 | -106.6815
25.7111 | -106.6942
25.7148 | -106.7063
25.7185 | -106.7192
25.7222 | -106.7337
25.7259 | -106.7478
25.7296 | -106.7603
25.7333 | -106.7723
25.737 | -106.7863
25.7407 | -106.7997
25.7444 | -106.8094
25.7481 | -106.8176
25.7519 | -106.8258
25.7556 | -106.8357
25.7593 | -106.845
25.763 | -106.852
25.7667 | -106.8598
25.7704 | -106.8696
25.7741 | -106.8788
25.7778 | -106.8873
25.7815 | -106.8953
25.7852 | -106.9015
25.7889 | -106.9066
25.7926 | -106.9108
25.7963 | -106.9142
25.8 | -106.9168
25.8037 | -106.9193
25.8074 | -106.9221
25.8111 | -106.9235
25.8148 | -106.9229
25.8185 | -106.9233
25.8222 | -106.9234
25.8259 | -106.9223
25.8296 | -106.9199
25.8333 | -106.9167
25.837 | -106.9141
25.8407 | -106.911
25.8444 | -106.9074
25.8481 | -106.9049
25.8519 | -106.9018
25.8556 | -106.8978
25.8593 | -106.8934
25.863 | -106.8898
25.8667 | -106.8857
25.8704 | -106.8783
25.8741 | -106.8706
25.8778 | -106.8636
25.8815 | -106.8574
25.8852 | -106.8508
25.8889 | -106.8436
25.8926 | -106.839
25.8963 | -106.8337
25.9 | -106.8281
25.9037 | -106.8217
25.9074 | -106.8143
25.9111 | -106.8081
25.9148 | -106.8028
25.9185 | -106.7995
25.9222 | -106.7966
25.9259 | -106.7922
25.9296 | -106.7884
25.9333 | -106.7862
25.937 | -106.7846
25.9407 | -106.7833
25.9444 | -106.7819
25.9481 | -106.7814
25.9519 | -106.782
25.9556 | -106.7822
25.9593 | -106.7798
25.963 | -106.778
25.9667 | -106.7787
25.9704 | -106.7798
25.9741 | -106.7823
25.9778 | -106.7844
25.9815 | -106.7862
25.9852 | -106.7877
25.9889 | -106.7882
25.9926 | -106.7914
25.9963 | -106.7953
26 | -106.7979
26.0037 | -106.8006
26.0074 | -106.803
26.0111 | -106.805
26.0148 | -106.8069
26.0185 | -106.8085
26.0222 | -106.8117
26.0259 | -106.8139
26.0296 | -106.8162
26.0333 | -106.8188
26.037 | -106.8204
26.0407 | -106.824
26.0444 | -106.8287
26.0481 | -106.8332
26.0519 | -106.8361
26.0556 | -106.8365
26.0593 | -106.8361
26.063 | -106.8379
26.0667 | -106.8398
26.0704 | -106.8419
26.0741 | -106.8427
26.0778 | -106.8434
26.0815 | -106.8453
26.0852 | -106.8469
26.0889 | -106.8477
26.0926 | -106.8486
26.0963 | -106.849
26.1 | -106.851
26.1037 | -106.8533
26.1074 | -106.8539
26.1111 | -106.8565
26.1148 | -106.8592
26.1185 | -106.8617
26.1222 | -106.8627
26.1259 | -106.8637
26.1296 | -106.8645
26.1333 | -106.8665
26.137 | -106.8672
26.1407 | -106.8677
26.1444 | -106.8689
26.1481 | -106.8718
26.1519 | -106.8742
26.1556 | -106.874
26.1593 | -106.8751
26.163 | -106.8761
26.1667 | -106.8775
26.1704 | -106.8797
26.1741 | -106.8818
26.1778 | -106.8844
26.1815 | -106.8862
26.1852 | -106.8877
26.1889 | -106.8887
26.1926 | -106.8887
26.1963 | -106.8903
26.2 | -106.8926
26.2037 | -106.8933
26.2074 | -106.8944
26.2111 | -106.8963
26.2148 | -106.8993
26.2185 | -106.9009
26.2222 | -106.9034
26.2259 | -106.9065
26.2296 | -106.91
26.2333 | -106.9139
26.237 | -106.917
26.2407 | -106.9207
26.2444 | -106.9243
26.2481 | -106.9288
26.2519 | -106.933
26.2556 | -106.9375
26.2593 | -106.9443
26.263 | -106.9492
26.2667 | -106.9547
26.2704 | -106.9611
26.2741 | -106.9679
26.2778 | -106.974
26.2815 | -106.9799
26.2852 | -106.9866
26.2889 | -106.9943
26.2926 | -107
26.2963 | -107.0059
26.3 | -107.0125
26.3037 | -107.0189
26.3074 | -107.0255
26.3111 | -107.0336
26.3148 | -107.042
26.3185 | -107.0494
26.3222 | -107.0583
26.3259 | -107.0676
26.3296 | -107.0756
26.3333 | -107.0851
26.337 | -107.0926
26.3407 | -107.0999
26.3444 | -107.1055
26.3481 | -107.1131
26.3519 | -107.1217
26.3556 | -107.1302
26.3593 | -107.1385
26.363 | -107.1447
26.3667 | -107.1507
26.3704 | -107.1565
26.3741 | -107.1625
26.3778 | -107.1676
26.3815 | -107.1733
26.3852 | -107.179
26.3889 | -107.1842
26.3926 | -107.1889
26.3963 | -107.1945
26.4 | -107.1993
26.4037 | -107.2052
26.4074 | -107.2112
26.4111 | -107.2157
26.4148 | -107.2196
26.4185 | -107.2227
26.4222 | -107.2255
26.4259 | -107.2288
26.4296 | -107.2321
26.4333 | -107.2344
26.437 | -107.2356
26.4407 | -107.2364
26.4444 | -107.2396
26.4481 | -107.2423
26.4519 | -107.243
26.4556 | -107.2426
26.4593 | -107.2427
26.463 | -107.2429
26.4667 | -107.243
26.4704 | -107.2419
26.4741 | -107.2401
26.4778 | -107.238
26.4815 | -107.2343
26.4852 | -107.2313
26.4889 | -107.2288
26.4926 | -107.2259
26.4963 | -107.2234
26.5 | -107.2225
26.5037 | -107.2206
26.5074 | -107.2179
26.5111 | -107.215
26.5148 | -107.2111
26.5185 | -107.2074
26.5222 | -107.2038
26.5259 | -107.2024
26.5296 | -107.2002
26.5333 | -107.198
26.537 | -107.1951
26.5407 | -107.1928
26.5444 | -107.1901
26.5481 | -107.1875
26.5519 | -107.1847
26.5556 | -107.1846
26.5593 | -107.1848
26.563 | -107.1848
26.5667 | -107.1843
26.5704 | -107.1836
26.5741 | -107.1836
26.5778 | -107.1855
26.5815 | -107.1866
26.5852 | -107.1858
26.5889 | -107.1858
26.5926 | -107.1884
26.5963 | -107.1907
26.6 | -107.193
26.6037 | -107.1947
26.6074 | -107.1974
26.6111 | -107.2019
26.6148 | -107.2066
26.6185 | -107.2114
26.6222 | -107.2141
26.6259 | -107.2185
26.6296 | -107.2227
26.6333 | -107.2263
26.637 | -107.2313
26.6407 | -107.2365
26.6444 | -107.2418
26.6481 | -107.248
26.6519 | -107.2547
26.6556 | -107.26
26.6593 | -107.2663
26.663 | -107.2717
26.6667 | -107.275
26.6704 | -107.2786
26.6741 | -107.2826
26.6778 | -107.2884
26.6815 | -107.294
26.6852 | -107.2964
26.6889 | -107.2991
26.6926 | -107.3011
26.6963 | -107.3012
26.7 | -107.3005
26.7037 | -107.3029
26.7074 | -107.3039
26.7111 | -107.3028
26.7148 | -107.2979
26.7185 | -107.2964
26.7222 | -107.2968
26.7259 | -107.2945
26.7296 | -107.2921
26.7333 | -107.2908
26.737 | -107.2903
26.7407 | -107.2894
26.7444 | -107.2857
26.7481 | -107.2813
26.7519 | -107.2774
26.7556 | -107.2751
26.7593 | -107.2694
26.763 | -107.2645
26.7667 | -107.2614
26.7704 | -107.2575
26.7741 | -107.2525
26.7778 | -107.2475
26.7815 | -107.2421
26.7852 | -107.2373
26.7889 | -107.2303
26.7926 | -107.2252
26.7963 | -107.2209
26.8 | -107.216
26.8037 | -107.211
26.8074 | -107.2055
26.8111 | -107.1994
26.8148 | -107.1948
26.8185 | -107.1889
26.8222 | -107.1832
26.8259 | -107.1774
26.8296 | -107.1729
26.8333 | -107.1691
26.837 | -107.1649
26.8407 | -107.1603
26.8444 | -107.1571
26.8481 | -107.1548
26.8519 | -107.1525
26.8556 | -107.15
26.8593 | -107.1474
26.863 | -107.1452
26.8667 | -107.1431
26.8704 | -107.1428
26.8741 | -107.1429
26.8778 | -107.1424
26.8815 | -107.1407
26.8852 | -107.1397
26.8889 | -107.1396
26.8926 | -107.141
26.8963 | -107.1431
26.9 | -107.147
26.9037 | -107.1498
26.9074 | -107.1534
26.9111 | -107.1565
26.9148 | -107.1604
26.9185 | -107.165
26.9222 | -107.1705
26.9259 | -107.1764
26.9296 | -107.1833
26.9333 | -107.1895
26.937 | -107.1953
26.9407 | -107.2003
26.9444 | -107.2069
26.9481 | -107.2148
26.9519 | -107.2218
26.9556 | -107.2279
26.9593 | -107.2354
26.963 | -107.2458
26.9667 | -107.2547
26.9704 | -107.2625
26.9741 | -107.2709
26.9778 | -107.2791
26.9815 | -107.2883
26.9852 | -107.297
26.9889 | -107.3053
26.9926 | -107.3137
26.9963 | -107.324
27 | -107.3339
27.0037 | -107.341
27.0074 | -107.3505
27.0111 | -107.3586
27.0148 | -107.3642
27.0185 | -107.3711
27.0222 | -107.3798
27.0259 | -107.3877
27.0296 | -107.3947
27.0333 | -107.4017
27.037 | -107.4074
27.0407 | -107.4136
27.0444 | -107.4225
27.0481 | -107.43
27.0519 | -107.4362
27.0556 | -107.4403
27.0593 | -107.4453
27.063 | -107.4516
27.0667 | -107.4581
27.0704 | -107.4645
27.0741 | -107.47
27.0778 | -107.4754
27.0815 | -107.4785
27.0852 | -107.4824
27.0889 | -107.4861
27.0926 | -107.4892
27.0963 | -107.4932
27.1 | -107.4969
27.1037 | -107.4998
27.1074 | -107.5028
27.1111 | -107.5022
27.1148 | -107.5006
27.1185 | -107.501
27.1222 | -107.5007
27.1259 | -107.5011
27.1296 | -107.5028
27.1333 | -107.503
27.137 | -107.502
27.1407 | -107.5032
27.1444 | -107.5032
27.1481 | -107.5005
27.1519 | -107.4986
27.1556 | -107.4968
27.1593 | -107.4936
27.163 | -107.4904
27.1667 | -107.4887
27.1704 | -107.4886
27.1741 | -107.4866
27.1778 | -107.4831
27.1815 | -107.4813
27.1852 | -107.4815
27.1889 | -107.4805
27.1926 | -107.4796
27.1963 | -107.4778
27.2 | -107.4766
27.2037 | -107.4729
27.2074 | -107.4698
27.2111 | -107.4671
27.2148 | -107.465
27.2185 | -107.4639
27.2222 | -107.4623
27.2259 | -107.4609
27.2296 | -107.4598
27.2333 | -107.4592
27.237 | -107.4576
27.2407 | -107.4565
27.2444 | -107.4543
27.2481 | -107.4531
27.2519 | -107.4507
27.2556 | -107.4474
27.2593 | -107.4439
27.263 | -107.4423
27.2667 | -107.4415
27.2704 | -107.4396
27.2741 | -107.4384
27.2778 | -107.4374
27.2815 | -107.4352
27.2852 | -107.4317
27.2889 | -107.4282
27.2926 | -107.4259
27.2963 | -107.4259
27.3 | -107.4261
27.3037 | -107.4254
27.3074 | -107.4252
27.3111 | -107.4247
27.3148 | -107.4221
27.3185 | -107.4187
27.3222 | -107.4166
27.3259 | -107.4159
27.3296 | -107.4151
27.3333 | -107.4132
27.337 | -107.4131
27.3407 | -107.4131
27.3444 | -107.4114
27.3481 | -107.4088
27.3519 | -107.4068
27.3556 | -107.4025
27.3593 | -107.3981
27.363 | -107.3945
27.3667 | -107.3914
27.3704 | -107.3894
27.3741 | -107.3871
27.3778 | -107.3844
27.3815 | -107.3801
27.3852 | -107.3765
27.3889 | -107.3731
27.3926 | -107.3684
27.3963 | -107.3643
27.4 | -107.3622
27.4037 | -107.3581
27.4074 | -107.3511
27.4111 | -107.3455
27.4148 | -107.3416
27.4185 | -107.3377
27.4222 | -107.3331
27.4259 | -107.328
27.4296 | -107.3231
27.4333 | -107.3178
27.437 | -107.3128
27.4407 | -107.3074
27.4444 | -107.2989
27.4481 | -107.2921
27.4519 | -107.2864
27.4556 | -107.2793
27.4593 | -107.2736
27.463 | -107.2673
27.4667 | -107.2599
27.4704 | -107.2547
27.4741 | -107.2506
27.4778 | -107.2449
27.4815 | -107.2385
27.4852 | -107.2341
27.4889 | -107.2287
27.4926 | -107.2239
27.4963 | -107.2213
27.5 | -107.2168
27.5037 | -107.2122
27.5074 | -107.2087
27.5111 | -107.206
27.5148 | -107.2024
27.5185 | -107.1989
27.5222 | -107.1953
27.5259 | -107.1923
27.5296 | -107.1893
27.5333 | -107.1874
27.537 | -107.1851
27.5407 | -107.1836
27.5444 | -107.1831
27.5481 | -107.1812
27.5519 | -107.1789
27.5556 | -107.1775
27.5593 | -107.1771
27.563 | -107.1766
27.5667 | -107.1744
27.5704 | -107.1711
27.5741 | -107.1696
27.5778 | -107.1693
27.5815 | -107.1694
27.5852 | -107.1694
27.5889 | -107.171
27.5926 | -107.1723
27.5963 | -107.172
27.6 | -107.1701
27.6037 | -107.1686
27.6074 | -107.168
27.6111 | -107.1681
27.6148 | -107.1677
27.6185 | -107.1667
27.6222 | -107.1659
27.6259 | -107.1659
27.6296 | -107.1656
27.6333 | -107.1647
27.637 | -107.1643
27.6407 | -107.1651
27.6444 | -107.1646
27.6481 | -107.1638
27.6519 | -107.1628
27.6556 | -107.1617
27.6593 | -107.1615
27.663 | -107.1613
27.6667 | -107.1612
27.6704 | -107.1606
27.6741 | -107.1595
27.6778 | -107.1591
27.6815 | -107.1579
27.6852 | -107.1553
27.6889 | -107.1527
27.6926 | -107.1496
27.6963 | -107.1463
27.7 | -107.144
27.7037 | -107.1414
27.7074 | -107.1372
27.7111 | -107.1339
27.7148 | -107.1309
27.7185 | -107.1286
27.7222 | -107.1252
27.7259 | -107.1208
27.7296 | -107.1175
27.7333 | -107.115
27.737 | -107.1111
27.7407 | -107.1066
27.7444 | -107.1033
27.7481 | -107.0998
27.7519 | -107.0969
27.7556 | -107.0938
27.7593 | -107.0895
27.763 | -107.0832
27.7667 | -107.0759
27.7704 | -107.0698
27.7741 | -107.0643
27.7778 | -107.0594
27.7815 | -107.0552
27.7852 | -107.0502
27.7889 | -107.0436
27.7926 | -107.0368
27.7963 | -107.0311
27.8 | -107.0276
27.8037 | -107.0227
27.8074 | -107.0166
27.8111 | -107.0118
27.8148 | -107.0077
27.8185 | -107.0027
27.8222 | -106.997
27.8259 | -106.9903
27.8296 | -106.984
27.8333 | -106.9798
27.837 | -106.9782
27.8407 | -106.9746
27.8444 | -106.9704
27.8481 | -106.9669
27.8519 | -106.9601
27.8556 | -106.9519
27.8593 | -106.9468
27.863 | -106.9406
27.8667 | -106.9361
27.8704 | -106.9326
27.8741 | -106.9283
27.8778 | -106.924
27.8815 | -106.9209
27.8852 | -106.9162
27.8889 | -106.9122
27.8926 | -106.9105
27.8963 | -106.9068
27.9 | -106.9029
27.9037 | -106.8983
27.9074 | -106.8954
27.9111 | -106.892
27.9148 | -106.8896
27.9185 | -106.8879
27.9222 | -106.8823
27.9259 | -106.8649
27.9296 | -106.8434
27.9333 | -106.8318
27.937 | -106.8233
27.9407 | -106.8161
27.9444 | -106.8084
27.9481 | -106.8007
27.9519 | -106.7919
27.9556 | -106.783
27.9593 | -106.7753
27.963 | -106.7666
27.9667 | -106.7579
27.9704 | -106.7479
27.9741 | -106.7369
27.9778 | -106.7259
27.9815 | -106.7144
27.9852 | -106.7033
27.9889 | -106.6929
27.9926 | -106.6809
27.9963 | -106.6673
28 | -106.6533
28.0037 | -106.6382