## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 287 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 292
Hood | 02-92B/P | 225
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1014.8 | hPa |  | PRC PRO 01
Humidity | 55.7 | %/Hr
Internal temperature of calibration bench | 24.28 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1021 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.1 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.98 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.01 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3072 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 11/09/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 13/09/2018
Technical Manager | Anthony HUE | Approval |  | 27/09/2018
CEO | Robert GLEMOT | Approval |  | 27/09/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4686 | 57.5717
0.012209 | -36.0318 | 38.4276
0.018313 | -34.9654 | 28.2696
0.024417 | -34.5029 | 22.0225
0.030521 | -34.3875 | 18.2506
0.036626 | -34.2413 | 15.1972
0.04273 | -34.1455 | 12.8345
0.048834 | -34.0999 | 11.3109
0.054938 | -34.0338 | 10.1667
0.061043 | -33.9722 | 9.0853
0.067147 | -33.9622 | 8.2649
0.073251 | -33.9631 | 7.506
0.079355 | -33.9813 | 6.8453
0.08546 | -33.9952 | 6.2502
0.091564 | -33.9903 | 5.8075
0.097668 | -33.9719 | 5.5587
0.10377 | -33.9492 | 5.3125
0.10988 | -33.9315 | 5.0358
0.11598 | -33.922 | 4.7433
0.12209 | -33.9235 | 4.4594
0.12819 | -33.9335 | 4.1758
0.13429 | -33.9444 | 3.8883
0.1404 | -33.9477 | 3.6102
0.1465 | -33.9495 | 3.3521
0.15261 | -33.9467 | 3.1309
0.15871 | -33.94 | 2.9672
0.16482 | -33.9317 | 2.8631
0.17092 | -33.9291 | 2.8238
0.17702 | -33.9267 | 2.7514
0.18313 | -33.9256 | 2.6734
0.18923 | -33.9249 | 2.5648
0.19534 | -33.9235 | 2.4384
0.20144 | -33.9219 | 2.2964
0.20754 | -33.9189 | 2.1755
0.21365 | -33.9178 | 2.0609
0.21975 | -33.9172 | 1.963
0.22586 | -33.9181 | 1.8756
0.23196 | -33.9205 | 1.8034
0.23807 | -33.9245 | 1.7409
0.24417 | -33.9286 | 1.683
0.25027 | -33.9325 | 1.6213
0.25638 | -33.9349 | 1.5604
0.26248 | -33.9367 | 1.4947
0.26859 | -33.9377 | 1.4306
0.27469 | -33.938 | 1.3606
0.2808 | -33.9391 | 1.2877
0.2869 | -33.9399 | 1.2174
0.293 | -33.9408 | 1.1508
0.29911 | -33.9417 | 1.091
0.30521 | -33.9427 | 1.0319
0.31132 | -33.9434 | 0.99273
0.31742 | -33.9446 | 0.94957
0.32353 | -33.9449 | 0.91411
0.32963 | -33.9457 | 0.88449
0.33573 | -33.9467 | 0.86352
0.34184 | -33.9479 | 0.84142
0.34794 | -33.95 | 0.81724
0.35405 | -33.952 | 0.79281
0.36015 | -33.9544 | 0.76713
0.36626 | -33.9563 | 0.73945
0.37236 | -33.9583 | 0.70268
0.37846 | -33.9604 | 0.6585
0.38457 | -33.9612 | 0.60163
0.39067 | -33.9622 | 0.55513
0.39678 | -33.9627 | 0.50342
0.40288 | -33.9631 | 0.44949
0.40899 | -33.963 | 0.39781
0.41509 | -33.9623 | 0.34716
0.42119 | -33.9622 | 0.30403
0.4273 | -33.9622 | 0.25802
0.4334 | -33.9621 | 0.21202
0.43951 | -33.9623 | 0.16882
0.44561 | -33.9632 | 0.12922
0.45172 | -33.9647 | 0.097679
0.45782 | -33.9655 | 0.06069
0.46392 | -33.9667 | 0.024674
0.47003 | -33.9677 | -0.0075536
0.47613 | -33.9685 | -0.041529
0.48224 | -33.969 | -0.076171
0.48834 | -33.968 | -0.11082
0.49445 | -33.9678 | -0.14599
0.50055 | -33.9671 | -0.17793
0.50665 | -33.9664 | -0.21204
0.51276 | -33.9653 | -0.24597
0.51886 | -33.9642 | -0.27984
0.52497 | -33.9629 | -0.30393
0.53107 | -33.9621 | -0.33327
0.53717 | -33.9613 | -0.36131
0.54328 | -33.9604 | -0.38408
0.54938 | -33.9602 | -0.40646
0.55549 | -33.9602 | -0.42736
0.56159 | -33.9612 | -0.43895
0.5677 | -33.9615 | -0.4535
0.5738 | -33.9625 | -0.4698
0.5799 | -33.9632 | -0.48477
0.58601 | -33.9643 | -0.49492
0.59211 | -33.9659 | -0.50146
0.59822 | -33.9673 | -0.5111
0.60432 | -33.9687 | -0.52337
0.61043 | -33.97 | -0.53149
0.61653 | -33.9712 | -0.54199
0.62263 | -33.9727 | -0.55155
0.62874 | -33.9734 | -0.56234
0.63484 | -33.9742 | -0.57173
0.64095 | -33.9747 | -0.58208
0.64705 | -33.975 | -0.59197
0.65316 | -33.9752 | -0.60456
0.65926 | -33.9753 | -0.62046
0.66536 | -33.9742 | -0.63786
0.67147 | -33.9738 | -0.65103
0.67757 | -33.9734 | -0.66535
0.68368 | -33.9726 | -0.68424
0.68978 | -33.9718 | -0.70133
0.69589 | -33.9709 | -0.72031
0.70199 | -33.9699 | -0.74299
0.70809 | -33.969 | -0.76209
0.7142 | -33.9683 | -0.7825
0.7203 | -33.9671 | -0.80161
0.72641 | -33.9665 | -0.8241
0.73251 | -33.9658 | -0.84407
0.73862 | -33.966 | -0.86587
0.74472 | -33.9657 | -0.8859
0.75082 | -33.9654 | -0.90806
0.75693 | -33.9657 | -0.92989
0.76303 | -33.9662 | -0.95191
0.76914 | -33.968 | -0.9725
0.77524 | -33.969 | -0.99213
0.78135 | -33.9702 | -1.0123
0.78745 | -33.9716 | -1.0322
0.79355 | -33.9728 | -1.0521
0.79966 | -33.9746 | -1.0715
0.80576 | -33.9768 | -1.0912
0.81187 | -33.9784 | -1.1129
0.81797 | -33.9804 | -1.1342
0.82408 | -33.982 | -1.1547
0.83018 | -33.9838 | -1.1753
0.83628 | -33.9857 | -1.1938
0.84239 | -33.9868 | -1.2142
0.84849 | -33.9882 | -1.2358
0.8546 | -33.9897 | -1.2556
0.8607 | -33.9909 | -1.2763
0.86681 | -33.992 | -1.2952
0.87291 | -33.9929 | -1.3139
0.87901 | -33.9934 | -1.3318
0.88512 | -33.9941 | -1.3498
0.89122 | -33.9948 | -1.368
0.89733 | -33.9952 | -1.3835
0.90343 | -33.9956 | -1.401
0.90953 | -33.9957 | -1.4152
0.91564 | -33.9957 | -1.4309
0.92174 | -33.9958 | -1.4458
0.92785 | -33.9959 | -1.4602
0.93395 | -33.9959 | -1.4739
0.94006 | -33.9961 | -1.4866
0.94616 | -33.9958 | -1.4983
0.95226 | -33.9958 | -1.5108
0.95837 | -33.9959 | -1.5245
0.96447 | -33.996 | -1.5372
0.97058 | -33.9958 | -1.5504
0.97668 | -33.9956 | -1.5636
0.98279 | -33.9953 | -1.5762
0.98889 | -33.995 | -1.5867
0.99499 | -33.995 | -1.6
1.0011 | -33.9948 | -1.6105
1.0072 | -33.9944 | -1.622
1.0133 | -33.994 | -1.6326
1.0194 | -33.9936 | -1.6454
1.0255 | -33.9933 | -1.659
1.0316 | -33.9927 | -1.6708
1.0377 | -33.9922 | -1.6843
1.0438 | -33.9916 | -1.6969
1.0499 | -33.9909 | -1.7121
1.056 | -33.9901 | -1.7263
1.0621 | -33.9895 | -1.7405
1.0682 | -33.989 | -1.7529
1.0743 | -33.9884 | -1.7683
1.0805 | -33.9878 | -1.7826
1.0866 | -33.9869 | -1.7976
1.0927 | -33.9861 | -1.8121
1.0988 | -33.9854 | -1.8269
1.1049 | -33.9846 | -1.841
1.111 | -33.9838 | -1.8564
1.1171 | -33.9831 | -1.8709
1.1232 | -33.9826 | -1.8861
1.1293 | -33.9822 | -1.9001
1.1354 | -33.9819 | -1.9151
1.1415 | -33.9815 | -1.9306
1.1476 | -33.9809 | -1.946
1.1537 | -33.9804 | -1.9607
1.1598 | -33.9805 | -1.9755
1.1659 | -33.9802 | -1.9889
1.172 | -33.98 | -2.0028
1.1781 | -33.9797 | -2.0167
1.1842 | -33.9795 | -2.0309
1.1903 | -33.9799 | -2.0418
1.1964 | -33.9797 | -2.0552
1.2025 | -33.9798 | -2.0691
1.2086 | -33.9801 | -2.0816
1.2147 | -33.9801 | -2.0941
1.2209 | -33.9805 | -2.1053
1.227 | -33.9811 | -2.116
1.2331 | -33.9815 | -2.1279
1.2392 | -33.9819 | -2.1396
1.2453 | -33.9823 | -2.1506
1.2514 | -33.9826 | -2.1605
1.2575 | -33.9832 | -2.1713
1.2636 | -33.984 | -2.18
1.2697 | -33.9846 | -2.1891
1.2758 | -33.9852 | -2.2001
1.2819 | -33.9859 | -2.2099
1.288 | -33.9865 | -2.2205
1.2941 | -33.9874 | -2.2289
1.3002 | -33.988 | -2.2367
1.3063 | -33.9885 | -2.2456
1.3124 | -33.9892 | -2.2544
1.3185 | -33.9899 | -2.2633
1.3246 | -33.9905 | -2.2722
1.3307 | -33.9914 | -2.2805
1.3368 | -33.9921 | -2.289
1.3429 | -33.9928 | -2.2977
1.349 | -33.9932 | -2.3068
1.3551 | -33.9939 | -2.3159
1.3613 | -33.9945 | -2.3238
1.3674 | -33.9951 | -2.3336
1.3735 | -33.9955 | -2.3431
1.3796 | -33.996 | -2.352
1.3857 | -33.9963 | -2.361
1.3918 | -33.9968 | -2.371
1.3979 | -33.9974 | -2.3799
1.404 | -33.9979 | -2.391
1.4101 | -33.9983 | -2.4012
1.4162 | -33.9986 | -2.4115
1.4223 | -33.9991 | -2.4227
1.4284 | -33.9996 | -2.4325
1.4345 | -33.9998 | -2.4439
1.4406 | -34.0001 | -2.4573
1.4467 | -34.0006 | -2.4687
1.4528 | -34.001 | -2.4799
1.4589 | -34.0015 | -2.4915
1.465 | -34.0017 | -2.5022
1.4711 | -34.0022 | -2.5135
1.4772 | -34.0027 | -2.5266
1.4833 | -34.0032 | -2.5395
1.4894 | -34.0037 | -2.5516
1.4955 | -34.0043 | -2.5634
1.5016 | -34.0048 | -2.5757
1.5078 | -34.0052 | -2.5895
1.5139 | -34.0056 | -2.6025
1.52 | -34.006 | -2.615
1.5261 | -34.0063 | -2.6275
1.5322 | -34.0066 | -2.6414
1.5383 | -34.007 | -2.6539
1.5444 | -34.0076 | -2.6665
1.5505 | -34.008 | -2.6798
1.5566 | -34.0085 | -2.6927
1.5627 | -34.009 | -2.7064
1.5688 | -34.0096 | -2.72
1.5749 | -34.0099 | -2.7326
1.581 | -34.0102 | -2.7459
1.5871 | -34.0106 | -2.7595
1.5932 | -34.0109 | -2.7723
1.5993 | -34.0113 | -2.7865
1.6054 | -34.0117 | -2.7997
1.6115 | -34.0121 | -2.8137
1.6176 | -34.0124 | -2.8263
1.6237 | -34.0126 | -2.8399
1.6298 | -34.0129 | -2.8524
1.6359 | -34.0132 | -2.8652
1.642 | -34.0136 | -2.8771
1.6482 | -34.0138 | -2.8901
1.6543 | -34.0141 | -2.9038
1.6604 | -34.0144 | -2.9158
1.6665 | -34.0145 | -2.9282
1.6726 | -34.0147 | -2.9407
1.6787 | -34.015 | -2.9534
1.6848 | -34.0149 | -2.966
1.6909 | -34.0152 | -2.9787
1.697 | -34.0154 | -2.9909
1.7031 | -34.0154 | -3.0035
1.7092 | -34.0154 | -3.0157
1.7153 | -34.0155 | -3.0277
1.7214 | -34.0153 | -3.0399
1.7275 | -34.0152 | -3.0525
1.7336 | -34.0152 | -3.0654
1.7397 | -34.0153 | -3.0777
1.7458 | -34.0153 | -3.09
1.7519 | -34.0153 | -3.1021
1.758 | -34.0153 | -3.1141
1.7641 | -34.0155 | -3.1257
1.7702 | -34.0154 | -3.1386
1.7763 | -34.0153 | -3.1502
1.7824 | -34.0153 | -3.1626
1.7885 | -34.0153 | -3.175
1.7947 | -34.0154 | -3.187
1.8008 | -34.0154 | -3.1983
1.8069 | -34.0153 | -3.2099
1.813 | -34.0154 | -3.2221
1.8191 | -34.0154 | -3.2348
1.8252 | -34.0155 | -3.2475
1.8313 | -34.0156 | -3.2596
1.8374 | -34.0157 | -3.2714
1.8435 | -34.0158 | -3.2833
1.8496 | -34.0158 | -3.2963
1.8557 | -34.016 | -3.309
1.8618 | -34.0162 | -3.3211
1.8679 | -34.0164 | -3.333
1.874 | -34.0165 | -3.3449
1.8801 | -34.0166 | -3.3572
1.8862 | -34.0169 | -3.3691
1.8923 | -34.017 | -3.3812
1.8984 | -34.0171 | -3.3931
1.9045 | -34.0172 | -3.4046
1.9106 | -34.0172 | -3.4166
1.9167 | -34.0173 | -3.4287
1.9228 | -34.0174 | -3.4405
1.9289 | -34.0175 | -3.4521
1.9351 | -34.0177 | -3.4628
1.9412 | -34.0179 | -3.474
1.9473 | -34.0182 | -3.4859
1.9534 | -34.0184 | -3.4972
1.9595 | -34.0186 | -3.5089
1.9656 | -34.0188 | -3.521
1.9717 | -34.0191 | -3.5326
1.9778 | -34.0194 | -3.5438
1.9839 | -34.0196 | -3.5551
1.99 | -34.0198 | -3.566
1.9961 | -34.0199 | -3.578
2.0022 | -34.0201 | -3.5892
2.0083 | -34.0202 | -3.6003
2.0144 | -34.0204 | -3.6116
2.0205 | -34.0208 | -3.6232
2.0266 | -34.0211 | -3.6344
2.0327 | -34.0213 | -3.6458
2.0388 | -34.0217 | -3.6572
2.0449 | -34.0219 | -3.6684
2.051 | -34.0222 | -3.6797
2.0571 | -34.0225 | -3.6909
2.0632 | -34.0227 | -3.7022
2.0693 | -34.0229 | -3.7142
2.0754 | -34.0232 | -3.7255
2.0816 | -34.0235 | -3.7375
2.0877 | -34.0236 | -3.7489
2.0938 | -34.024 | -3.7608
2.0999 | -34.0242 | -3.7724
2.106 | -34.0244 | -3.7839
2.1121 | -34.0245 | -3.7958
2.1182 | -34.0248 | -3.8064
2.1243 | -34.025 | -3.8171
2.1304 | -34.0252 | -3.8281
2.1365 | -34.0255 | -3.8397
2.1426 | -34.0257 | -3.8515
2.1487 | -34.026 | -3.8627
2.1548 | -34.0264 | -3.8744
2.1609 | -34.0268 | -3.8857
2.167 | -34.0272 | -3.8968
2.1731 | -34.0275 | -3.9087
2.1792 | -34.0277 | -3.9196
2.1853 | -34.028 | -3.9314
2.1914 | -34.0283 | -3.9426
2.1975 | -34.0287 | -3.954
2.2036 | -34.0289 | -3.9659
2.2097 | -34.0291 | -3.9772
2.2158 | -34.0294 | -3.9877
2.222 | -34.0297 | -3.9981
2.2281 | -34.0299 | -4.0093
2.2342 | -34.0302 | -4.0205
2.2403 | -34.0304 | -4.031
2.2464 | -34.0306 | -4.0417
2.2525 | -34.0308 | -4.053
2.2586 | -34.0312 | -4.0635
2.2647 | -34.0314 | -4.0738
2.2708 | -34.0317 | -4.0849
2.2769 | -34.032 | -4.0954
2.283 | -34.0322 | -4.1065
2.2891 | -34.0324 | -4.1173
2.2952 | -34.0325 | -4.1277
2.3013 | -34.0327 | -4.1385
2.3074 | -34.033 | -4.1487
2.3135 | -34.0332 | -4.16
2.3196 | -34.0333 | -4.17
2.3257 | -34.0335 | -4.1805
2.3318 | -34.0337 | -4.1914
2.3379 | -34.0338 | -4.2021
2.344 | -34.0339 | -4.2133
2.3501 | -34.0341 | -4.2239
2.3562 | -34.0343 | -4.2342
2.3623 | -34.0344 | -4.2445
2.3685 | -34.0344 | -4.2549
2.3746 | -34.0346 | -4.266
2.3807 | -34.0347 | -4.2763
2.3868 | -34.0349 | -4.2869
2.3929 | -34.0351 | -4.2974
2.399 | -34.0351 | -4.3078
2.4051 | -34.0354 | -4.3182
2.4112 | -34.0355 | -4.3287
2.4173 | -34.0357 | -4.3392
2.4234 | -34.0357 | -4.3498
2.4295 | -34.0359 | -4.3606
2.4356 | -34.036 | -4.3713
2.4417 | -34.0361 | -4.3823
2.4478 | -34.0362 | -4.393
2.4539 | -34.0364 | -4.4031
2.46 | -34.0364 | -4.4145
2.4661 | -34.0366 | -4.4255
2.4722 | -34.0367 | -4.4355
2.4783 | -34.0368 | -4.4459
2.4844 | -34.0369 | -4.4568
2.4905 | -34.0369 | -4.4676
2.4966 | -34.0371 | -4.4782
2.5027 | -34.0372 | -4.4885
2.5089 | -34.0373 | -4.4993
2.515 | -34.0373 | -4.51
2.5211 | -34.0374 | -4.5206
2.5272 | -34.0375 | -4.5315
2.5333 | -34.0376 | -4.5423
2.5394 | -34.0377 | -4.553
2.5455 | -34.0378 | -4.5642
2.5516 | -34.038 | -4.5744
2.5577 | -34.0382 | -4.5858
2.5638 | -34.0383 | -4.5974
2.5699 | -34.0383 | -4.6087
2.576 | -34.0384 | -4.6202
2.5821 | -34.0385 | -4.6314
2.5882 | -34.0386 | -4.6424
2.5943 | -34.0387 | -4.6531
2.6004 | -34.0389 | -4.664
2.6065 | -34.039 | -4.6743
2.6126 | -34.0391 | -4.6847
2.6187 | -34.0392 | -4.6951
2.6248 | -34.0393 | -4.7058
2.6309 | -34.0395 | -4.7172
2.637 | -34.0398 | -4.7287
2.6431 | -34.0399 | -4.7391
2.6492 | -34.04 | -4.7504
2.6554 | -34.0402 | -4.7611
2.6615 | -34.0403 | -4.7721
2.6676 | -34.0405 | -4.7834
2.6737 | -34.0408 | -4.7945
2.6798 | -34.0411 | -4.8055
2.6859 | -34.0413 | -4.8167
2.692 | -34.0414 | -4.8275
2.6981 | -34.0416 | -4.8386
2.7042 | -34.0418 | -4.8496
2.7103 | -34.042 | -4.8601
2.7164 | -34.0423 | -4.871
2.7225 | -34.0426 | -4.8823
2.7286 | -34.0428 | -4.8933
2.7347 | -34.043 | -4.9037
2.7408 | -34.0433 | -4.9137
2.7469 | -34.0435 | -4.9249
2.753 | -34.0438 | -4.9354
2.7591 | -34.0441 | -4.9466
2.7652 | -34.0443 | -4.9581
2.7713 | -34.0444 | -4.9688
2.7774 | -34.0447 | -4.9796
2.7835 | -34.0449 | -4.9905
2.7896 | -34.0451 | -5.001
2.7958 | -34.0454 | -5.0116
2.8019 | -34.0457 | -5.022
2.808 | -34.0459 | -5.0327
2.8141 | -34.0462 | -5.0438
2.8202 | -34.0465 | -5.0546
2.8263 | -34.0468 | -5.0654
2.8324 | -34.0472 | -5.0768
2.8385 | -34.0476 | -5.0875
2.8446 | -34.0479 | -5.0982
2.8507 | -34.0483 | -5.1089
2.8568 | -34.0486 | -5.1193
2.8629 | -34.0489 | -5.1298
2.869 | -34.0492 | -5.1408
2.8751 | -34.0495 | -5.1515
2.8812 | -34.0499 | -5.1627
2.8873 | -34.0502 | -5.1735
2.8934 | -34.0505 | -5.1845
2.8995 | -34.0508 | -5.1955
2.9056 | -34.0511 | -5.2065
2.9117 | -34.0514 | -5.2175
2.9178 | -34.0518 | -5.2281
2.9239 | -34.0522 | -5.2396
2.93 | -34.0526 | -5.2509
2.9361 | -34.0529 | -5.2616
2.9423 | -34.0532 | -5.2722
2.9484 | -34.0536 | -5.2831
2.9545 | -34.054 | -5.294
2.9606 | -34.0544 | -5.3053
2.9667 | -34.0547 | -5.3163
2.9728 | -34.055 | -5.3275
2.9789 | -34.0555 | -5.3385
2.985 | -34.0559 | -5.3497
2.9911 | -34.0562 | -5.3608
2.9972 | -34.0567 | -5.3718
3.0033 | -34.0571 | -5.3832
3.0094 | -34.0574 | -5.3942
3.0155 | -34.0579 | -5.405
3.0216 | -34.0582 | -5.4161
3.0277 | -34.0585 | -5.4277
3.0338 | -34.0589 | -5.439
3.0399 | -34.0593 | -5.4502
3.046 | -34.0596 | -5.4616
3.0521 | -34.0599 | -5.4726
3.0582 | -34.0602 | -5.4844
3.0643 | -34.0606 | -5.4958
3.0704 | -34.0609 | -5.5074
3.0765 | -34.0611 | -5.5184
3.0827 | -34.0614 | -5.5295
3.0888 | -34.0617 | -5.5406
3.0949 | -34.0621 | -5.5518
3.101 | -34.0624 | -5.5631
3.1071 | -34.0627 | -5.5747
3.1132 | -34.0631 | -5.5863
3.1193 | -34.0634 | -5.5977
3.1254 | -34.0638 | -5.6089
3.1315 | -34.0642 | -5.6206
3.1376 | -34.0644 | -5.6319
3.1437 | -34.0648 | -5.6433
3.1498 | -34.0651 | -5.6552
3.1559 | -34.0654 | -5.6665
3.162 | -34.0656 | -5.6784
3.1681 | -34.0659 | -5.6902
3.1742 | -34.0662 | -5.7016
3.1803 | -34.0665 | -5.7131
3.1864 | -34.0668 | -5.7242
3.1925 | -34.0671 | -5.7356
3.1986 | -34.0674 | -5.7474
3.2047 | -34.0676 | -5.7592
3.2108 | -34.0679 | -5.7702
3.2169 | -34.0682 | -5.7817
3.223 | -34.0684 | -5.7932
3.2292 | -34.0686 | -5.8049
3.2353 | -34.0688 | -5.8156
3.2414 | -34.069 | -5.8268
3.2475 | -34.0692 | -5.8381
3.2536 | -34.0695 | -5.8494
3.2597 | -34.0697 | -5.8602
3.2658 | -34.0699 | -5.871
3.2719 | -34.0701 | -5.8819
3.278 | -34.0704 | -5.8933
3.2841 | -34.0706 | -5.9042
3.2902 | -34.0708 | -5.9155
3.2963 | -34.0711 | -5.9268
3.3024 | -34.0712 | -5.9378
3.3085 | -34.0715 | -5.9492
3.3146 | -34.0716 | -5.9602
3.3207 | -34.0718 | -5.9716
3.3268 | -34.072 | -5.9829
3.3329 | -34.0722 | -5.9934
3.339 | -34.0724 | -6.0041
3.3451 | -34.0725 | -6.0151
3.3512 | -34.0727 | -6.0257
3.3573 | -34.0729 | -6.0368
3.3634 | -34.0731 | -6.0478
3.3696 | -34.0733 | -6.059
3.3757 | -34.0734 | -6.0696
3.3818 | -34.0736 | -6.0804
3.3879 | -34.0738 | -6.0913
3.394 | -34.0739 | -6.1026
3.4001 | -34.0741 | -6.1133
3.4062 | -34.0742 | -6.124
3.4123 | -34.0743 | -6.1345
3.4184 | -34.0745 | -6.1457
3.4245 | -34.0747 | -6.1563
3.4306 | -34.0748 | -6.1669
3.4367 | -34.075 | -6.1773
3.4428 | -34.0752 | -6.1876
3.4489 | -34.0753 | -6.1984
3.455 | -34.0754 | -6.2091
3.4611 | -34.0756 | -6.2203
3.4672 | -34.0757 | -6.2312
3.4733 | -34.0759 | -6.2422
3.4794 | -34.0761 | -6.2529
3.4855 | -34.0762 | -6.2632
3.4916 | -34.0764 | -6.2735
3.4977 | -34.0765 | -6.2847
3.5038 | -34.0767 | -6.2954
3.5099 | -34.0768 | -6.3056
3.5161 | -34.077 | -6.3159
3.5222 | -34.0771 | -6.3262
3.5283 | -34.0772 | -6.3369
3.5344 | -34.0773 | -6.3476
3.5405 | -34.0775 | -6.3581
3.5466 | -34.0776 | -6.3689
3.5527 | -34.0777 | -6.3797
3.5588 | -34.0778 | -6.3902
3.5649 | -34.0779 | -6.4008
3.571 | -34.078 | -6.4114
3.5771 | -34.0782 | -6.4224
3.5832 | -34.0783 | -6.4322
3.5893 | -34.0785 | -6.4427
3.5954 | -34.0787 | -6.4534
3.6015 | -34.0789 | -6.4643
3.6076 | -34.0791 | -6.4747
3.6137 | -34.0792 | -6.4847
3.6198 | -34.0795 | -6.495
3.6259 | -34.0796 | -6.5058
3.632 | -34.0798 | -6.5164
3.6381 | -34.0799 | -6.5271
3.6442 | -34.0801 | -6.5376
3.6503 | -34.0803 | -6.5477
3.6565 | -34.0805 | -6.5584
3.6626 | -34.0806 | -6.5689
3.6687 | -34.0809 | -6.5795
3.6748 | -34.0811 | -6.5898
3.6809 | -34.0814 | -6.6004
3.687 | -34.0815 | -6.6113
3.6931 | -34.0818 | -6.6215
3.6992 | -34.082 | -6.6321
3.7053 | -34.0822 | -6.6426
3.7114 | -34.0823 | -6.6532
3.7175 | -34.0825 | -6.6639
3.7236 | -34.0828 | -6.6739
3.7297 | -34.083 | -6.6846
3.7358 | -34.0832 | -6.6955
3.7419 | -34.0834 | -6.7062
3.748 | -34.0836 | -6.7167
3.7541 | -34.0838 | -6.7273
3.7602 | -34.0841 | -6.7378
3.7663 | -34.0843 | -6.7482
3.7724 | -34.0845 | -6.7588
3.7785 | -34.0848 | -6.7692
3.7846 | -34.085 | -6.7796
3.7907 | -34.0853 | -6.7905
3.7969 | -34.0855 | -6.8011
3.803 | -34.0858 | -6.8113
3.8091 | -34.086 | -6.8218
3.8152 | -34.0863 | -6.8323
3.8213 | -34.0865 | -6.8431
3.8274 | -34.0867 | -6.8534
3.8335 | -34.0871 | -6.8641
3.8396 | -34.0874 | -6.8747
3.8457 | -34.0877 | -6.8854
3.8518 | -34.088 | -6.8958
3.8579 | -34.0883 | -6.9064
3.864 | -34.0886 | -6.9171
3.8701 | -34.0889 | -6.9281
3.8762 | -34.0891 | -6.9385
3.8823 | -34.0893 | -6.9493
3.8884 | -34.0897 | -6.9599
3.8945 | -34.09 | -6.9707
3.9006 | -34.0903 | -6.9816
3.9067 | -34.0905 | -6.9927
3.9128 | -34.0908 | -7.0034
3.9189 | -34.0911 | -7.0138
3.925 | -34.0913 | -7.0246
3.9311 | -34.0916 | -7.0348
3.9372 | -34.0919 | -7.0456
3.9434 | -34.0922 | -7.0562
3.9495 | -34.0925 | -7.067
3.9556 | -34.0928 | -7.0774
3.9617 | -34.0931 | -7.0876
3.9678 | -34.0934 | -7.098
3.9739 | -34.0937 | -7.1086
3.98 | -34.0939 | -7.1193
3.9861 | -34.0942 | -7.1299
3.9922 | -34.0944 | -7.1405
3.9983 | -34.0947 | -7.1508
4.0044 | -34.095 | -7.1609
4.0105 | -34.0952 | -7.1716
4.0166 | -34.0956 | -7.1823
4.0227 | -34.0958 | -7.1928
4.0288 | -34.0961 | -7.2038
4.0349 | -34.0963 | -7.2142
4.041 | -34.0965 | -7.2247
4.0471 | -34.0967 | -7.2354
4.0532 | -34.097 | -7.2457
4.0593 | -34.0973 | -7.2564
4.0654 | -34.0976 | -7.2668
4.0715 | -34.0979 | -7.2778
4.0776 | -34.0982 | -7.2883
4.0838 | -34.0984 | -7.2994
4.0899 | -34.0987 | -7.3097
4.096 | -34.099 | -7.3202
4.1021 | -34.0992 | -7.3306
4.1082 | -34.0995 | -7.3409
4.1143 | -34.0997 | -7.3511
4.1204 | -34.1 | -7.3613
4.1265 | -34.1002 | -7.3716
4.1326 | -34.1005 | -7.3819
4.1387 | -34.1008 | -7.3923
4.1448 | -34.101 | -7.4029
4.1509 | -34.1012 | -7.4134
4.157 | -34.1015 | -7.4243
4.1631 | -34.1017 | -7.435
4.1692 | -34.102 | -7.4454
4.1753 | -34.1022 | -7.4559
4.1814 | -34.1024 | -7.466
4.1875 | -34.1027 | -7.4767
4.1936 | -34.1029 | -7.4873
4.1997 | -34.1031 | -7.4974
4.2058 | -34.1034 | -7.5082
4.2119 | -34.1036 | -7.5185
4.218 | -34.1038 | -7.5286
4.2241 | -34.104 | -7.5389
4.2303 | -34.1043 | -7.5492
4.2364 | -34.1045 | -7.5596
4.2425 | -34.1047 | -7.5701
4.2486 | -34.1049 | -7.5806
4.2547 | -34.105 | -7.5908
4.2608 | -34.1052 | -7.6013
4.2669 | -34.1054 | -7.6114
4.273 | -34.1056 | -7.6219
4.2791 | -34.1058 | -7.6323
4.2852 | -34.106 | -7.6429
4.2913 | -34.1062 | -7.653
4.2974 | -34.1064 | -7.6637
4.3035 | -34.1066 | -7.6743
4.3096 | -34.1068 | -7.6849
4.3157 | -34.107 | -7.6951
4.3218 | -34.1072 | -7.7053
4.3279 | -34.1074 | -7.7157
4.334 | -34.1076 | -7.726
4.3401 | -34.1078 | -7.7359
4.3462 | -34.108 | -7.7464
4.3523 | -34.1083 | -7.7563
4.3584 | -34.1084 | -7.7663
4.3645 | -34.1086 | -7.7765
4.3707 | -34.1088 | -7.7866
4.3768 | -34.109 | -7.7968
4.3829 | -34.1092 | -7.8074
4.389 | -34.1094 | -7.8177
4.3951 | -34.1096 | -7.8277
4.4012 | -34.1098 | -7.8383
4.4073 | -34.1099 | -7.8484
4.4134 | -34.1101 | -7.8586
4.4195 | -34.1104 | -7.8688
4.4256 | -34.1106 | -7.8789
4.4317 | -34.1108 | -7.8891
4.4378 | -34.111 | -7.899
4.4439 | -34.1112 | -7.9089
4.45 | -34.1113 | -7.9191
4.4561 | -34.1115 | -7.9293
4.4622 | -34.1116 | -7.9395
4.4683 | -34.1118 | -7.9497
4.4744 | -34.112 | -7.9602
4.4805 | -34.1122 | -7.97
4.4866 | -34.1124 | -7.9802
4.4927 | -34.1127 | -7.9906
4.4988 | -34.1128 | -8.0007
4.5049 | -34.1131 | -8.0105
4.511 | -34.1133 | -8.0204
4.5172 | -34.1135 | -8.0306
4.5233 | -34.1137 | -8.0411
4.5294 | -34.1139 | -8.0511
4.5355 | -34.1141 | -8.0613
4.5416 | -34.1143 | -8.0714
4.5477 | -34.1145 | -8.0816
4.5538 | -34.1147 | -8.0913
4.5599 | -34.1149 | -8.1014
4.566 | -34.1151 | -8.1116
4.5721 | -34.1153 | -8.1217
4.5782 | -34.1156 | -8.1321
4.5843 | -34.1158 | -8.1423
4.5904 | -34.116 | -8.1526
4.5965 | -34.1163 | -8.1625
4.6026 | -34.1165 | -8.1731
4.6087 | -34.1167 | -8.1832
4.6148 | -34.1169 | -8.1937
4.6209 | -34.1171 | -8.2037
4.627 | -34.1174 | -8.2141
4.6331 | -34.1175 | -8.2241
4.6392 | -34.1177 | -8.2343
4.6453 | -34.1179 | -8.2445
4.6514 | -34.1182 | -8.2547
4.6576 | -34.1184 | -8.2652
4.6637 | -34.1185 | -8.275
4.6698 | -34.1187 | -8.285
4.6759 | -34.119 | -8.2956
4.682 | -34.1191 | -8.3057
4.6881 | -34.1193 | -8.3158
4.6942 | -34.1195 | -8.3259
4.7003 | -34.1197 | -8.3361
4.7064 | -34.12 | -8.3462
4.7125 | -34.1202 | -8.3567
4.7186 | -34.1204 | -8.3666
4.7247 | -34.1207 | -8.3769
4.7308 | -34.1209 | -8.387
4.7369 | -34.1211 | -8.3973
4.743 | -34.1213 | -8.4079
4.7491 | -34.1216 | -8.4184
4.7552 | -34.1218 | -8.4287
4.7613 | -34.122 | -8.4391
4.7674 | -34.1222 | -8.4491
4.7735 | -34.1224 | -8.4598
4.7796 | -34.1226 | -8.4702
4.7857 | -34.1229 | -8.4806
4.7918 | -34.1231 | -8.4906
4.7979 | -34.1234 | -8.5007
4.8041 | -34.1236 | -8.511
4.8102 | -34.1238 | -8.5209
4.8163 | -34.1241 | -8.5311
4.8224 | -34.1243 | -8.5414
4.8285 | -34.1245 | -8.5513
4.8346 | -34.1248 | -8.5613
4.8407 | -34.125 | -8.5717
4.8468 | -34.1253 | -8.5822
4.8529 | -34.1255 | -8.5927
4.859 | -34.1257 | -8.6033
4.8651 | -34.126 | -8.6134
4.8712 | -34.1262 | -8.6234
4.8773 | -34.1264 | -8.6338
4.8834 | -34.1267 | -8.644
4.8895 | -34.1269 | -8.6541
4.8956 | -34.1271 | -8.6643
4.9017 | -34.1273 | -8.6745
4.9078 | -34.1275 | -8.685
4.9139 | -34.1278 | -8.6949
4.92 | -34.1281 | -8.7049
4.9261 | -34.1283 | -8.7148
4.9322 | -34.1286 | -8.7251
4.9383 | -34.1289 | -8.7353
4.9445 | -34.1291 | -8.7455
4.9506 | -34.1293 | -8.7556
4.9567 | -34.1296 | -8.7657
4.9628 | -34.1298 | -8.7759
4.9689 | -34.1301 | -8.7865
4.975 | -34.1303 | -8.7971
4.9811 | -34.1305 | -8.807
4.9872 | -34.1308 | -8.8171
4.9933 | -34.1312 | -8.8277
4.9994 | -34.1314 | -8.8383
5.0055 | -34.1317 | -8.8486
5.0116 | -34.132 | -8.8593
5.0177 | -34.1323 | -8.8698
5.0238 | -34.1326 | -8.8804
5.0299 | -34.1329 | -8.8911
5.036 | -34.1332 | -8.9013
5.0421 | -34.1334 | -8.9115
5.0482 | -34.1337 | -8.9218
5.0543 | -34.134 | -8.9318
5.0604 | -34.1343 | -8.9421
5.0665 | -34.1346 | -8.9522
5.0726 | -34.1348 | -8.9628
5.0787 | -34.1351 | -8.9733
5.0848 | -34.1354 | -8.9837
5.091 | -34.1357 | -8.9938
5.0971 | -34.136 | -9.0038
5.1032 | -34.1362 | -9.0144
5.1093 | -34.1365 | -9.0248
5.1154 | -34.1368 | -9.0348
5.1215 | -34.137 | -9.045
5.1276 | -34.1373 | -9.0551
5.1337 | -34.1377 | -9.0651
5.1398 | -34.1379 | -9.0752
5.1459 | -34.1382 | -9.0858
5.152 | -34.1385 | -9.0962
5.1581 | -34.1388 | -9.1067
5.1642 | -34.1391 | -9.1171
5.1703 | -34.1393 | -9.1273
5.1764 | -34.1397 | -9.1374
5.1825 | -34.1399 | -9.1475
5.1886 | -34.1402 | -9.1574
5.1947 | -34.1405 | -9.1677
5.2008 | -34.1408 | -9.1781
5.2069 | -34.1411 | -9.1886
5.213 | -34.1414 | -9.1987
5.2191 | -34.1417 | -9.209
5.2252 | -34.142 | -9.2194
5.2314 | -34.1422 | -9.2299
5.2375 | -34.1425 | -9.2399
5.2436 | -34.1428 | -9.2502
5.2497 | -34.1431 | -9.2602
5.2558 | -34.1434 | -9.2709
5.2619 | -34.1437 | -9.2814
5.268 | -34.144 | -9.2919
5.2741 | -34.1443 | -9.3023
5.2802 | -34.1446 | -9.3129
5.2863 | -34.145 | -9.3232
5.2924 | -34.1453 | -9.3337
5.2985 | -34.1456 | -9.344
5.3046 | -34.1459 | -9.3544
5.3107 | -34.1462 | -9.3649
5.3168 | -34.1465 | -9.3753
5.3229 | -34.1467 | -9.3856
5.329 | -34.147 | -9.396
5.3351 | -34.1474 | -9.406
5.3412 | -34.1476 | -9.4163
5.3473 | -34.1479 | -9.4266
5.3534 | -34.1483 | -9.4371
5.3595 | -34.1485 | -9.4474
5.3656 | -34.1489 | -9.4577
5.3717 | -34.1492 | -9.4683
5.3779 | -34.1494 | -9.479
5.384 | -34.1497 | -9.4893
5.3901 | -34.1501 | -9.4999
5.3962 | -34.1504 | -9.5101
5.4023 | -34.1507 | -9.5206
5.4084 | -34.151 | -9.5309
5.4145 | -34.1513 | -9.541
5.4206 | -34.1516 | -9.5514
5.4267 | -34.1519 | -9.562
5.4328 | -34.1522 | -9.5724
5.4389 | -34.1525 | -9.5825
5.445 | -34.1528 | -9.5931
5.4511 | -34.1531 | -9.6033
5.4572 | -34.1535 | -9.6138
5.4633 | -34.1538 | -9.6241
5.4694 | -34.154 | -9.6345
5.4755 | -34.1543 | -9.6446
5.4816 | -34.1546 | -9.655
5.4877 | -34.155 | -9.6652
5.4938 | -34.1552 | -9.6752
5.4999 | -34.1556 | -9.6854
5.506 | -34.1559 | -9.6957
5.5121 | -34.1562 | -9.706
5.5183 | -34.1565 | -9.716
5.5244 | -34.1568 | -9.7265
5.5305 | -34.1571 | -9.7367
5.5366 | -34.1574 | -9.7468
5.5427 | -34.1576 | -9.7573
5.5488 | -34.1579 | -9.7676
5.5549 | -34.1582 | -9.7779
5.561 | -34.1585 | -9.7885
5.5671 | -34.1589 | -9.7986
5.5732 | -34.1592 | -9.8092
5.5793 | -34.1595 | -9.8198
5.5854 | -34.1598 | -9.83
5.5915 | -34.1601 | -9.8406
5.5976 | -34.1604 | -9.851
5.6037 | -34.1607 | -9.8614
5.6098 | -34.1609 | -9.8717
5.6159 | -34.1613 | -9.8818
5.622 | -34.1616 | -9.8922
5.6281 | -34.1619 | -9.9025
5.6342 | -34.1622 | -9.9132
5.6403 | -34.1625 | -9.9236
5.6464 | -34.1628 | -9.934
5.6525 | -34.1631 | -9.9445
5.6586 | -34.1634 | -9.9545
5.6648 | -34.1637 | -9.9649
5.6709 | -34.164 | -9.9754
5.677 | -34.1643 | -9.9857
5.6831 | -34.1646 | -9.9961
5.6892 | -34.1649 | -10.0066
5.6953 | -34.1652 | -10.0169
5.7014 | -34.1655 | -10.0273
5.7075 | -34.1658 | -10.0376
5.7136 | -34.1661 | -10.0479
5.7197 | -34.1663 | -10.0582
5.7258 | -34.1666 | -10.0683
5.7319 | -34.1669 | -10.0787
5.738 | -34.1672 | -10.089
5.7441 | -34.1676 | -10.0992
5.7502 | -34.1679 | -10.1093
5.7563 | -34.1682 | -10.1196
5.7624 | -34.1685 | -10.1301
5.7685 | -34.1688 | -10.1405
5.7746 | -34.1691 | -10.1511
5.7807 | -34.1693 | -10.1615
5.7868 | -34.1696 | -10.172
5.7929 | -34.1699 | -10.1822
5.799 | -34.1702 | -10.1923
5.8052 | -34.1705 | -10.2025
5.8113 | -34.1708 | -10.2127
5.8174 | -34.171 | -10.2229
5.8235 | -34.1714 | -10.2331
5.8296 | -34.1716 | -10.2436
5.8357 | -34.1719 | -10.2538
5.8418 | -34.1722 | -10.2642
5.8479 | -34.1725 | -10.2746
5.854 | -34.1728 | -10.2851
5.8601 | -34.1731 | -10.2952
5.8662 | -34.1733 | -10.3057
5.8723 | -34.1736 | -10.3158
5.8784 | -34.1739 | -10.326
5.8845 | -34.1742 | -10.3362
5.8906 | -34.1745 | -10.3466
5.8967 | -34.1748 | -10.3567
5.9028 | -34.1751 | -10.3668
5.9089 | -34.1755 | -10.3769
5.915 | -34.1757 | -10.387
5.9211 | -34.176 | -10.3972
5.9272 | -34.1763 | -10.4072
5.9333 | -34.1766 | -10.4173
5.9394 | -34.1769 | -10.4275
5.9455 | -34.1772 | -10.4376
5.9517 | -34.1776 | -10.4477
5.9578 | -34.1779 | -10.4581
5.9639 | -34.1782 | -10.468
5.97 | -34.1785 | -10.4783
5.9761 | -34.1789 | -10.4884
5.9822 | -34.1791 | -10.4988
5.9883 | -34.1794 | -10.5091
5.9944 | -34.1798 | -10.5194
6.0005 | -34.1801 | -10.5298
6.0066 | -34.1804 | -10.5398
6.0127 | -34.1807 | -10.5498
6.0188 | -34.181 | -10.5602
6.0249 | -34.1813 | -10.5705
6.031 | -34.1816 | -10.5805
6.0371 | -34.1819 | -10.5908
6.0432 | -34.1822 | -10.6009
6.0493 | -34.1825 | -10.611
6.0554 | -34.1828 | -10.621
6.0615 | -34.1831 | -10.6313
6.0676 | -34.1834 | -10.6416
6.0737 | -34.1838 | -10.652
6.0798 | -34.184 | -10.6622
6.0859 | -34.1844 | -10.6723
6.0921 | -34.1847 | -10.6823
6.0982 | -34.185 | -10.6925
6.1043 | -34.1853 | -10.7025
6.1104 | -34.1856 | -10.7129
6.1165 | -34.186 | -10.7227
6.1226 | -34.1863 | -10.7328
6.1287 | -34.1866 | -10.743
6.1348 | -34.1869 | -10.7531
6.1409 | -34.1872 | -10.7632
6.147 | -34.1875 | -10.7733
6.1531 | -34.1878 | -10.7835
6.1592 | -34.1881 | -10.7935
6.1653 | -34.1884 | -10.8036
6.1714 | -34.1887 | -10.8138
6.1775 | -34.189 | -10.8241
6.1836 | -34.1893 | -10.8342
6.1897 | -34.1897 | -10.8444
6.1958 | -34.19 | -10.8548
6.2019 | -34.1903 | -10.865
6.208 | -34.1906 | -10.8752
6.2141 | -34.1909 | -10.8853
6.2202 | -34.1913 | -10.8956
6.2263 | -34.1916 | -10.9057
6.2325 | -34.1919 | -10.9156
6.2386 | -34.1922 | -10.9254
6.2447 | -34.1925 | -10.9358
6.2508 | -34.1929 | -10.946
6.2569 | -34.1932 | -10.956
6.263 | -34.1935 | -10.9662
6.2691 | -34.1938 | -10.9762
6.2752 | -34.1941 | -10.9864
6.2813 | -34.1944 | -10.9965
6.2874 | -34.1947 | -11.0065
6.2935 | -34.195 | -11.0166
6.2996 | -34.1953 | -11.0267
6.3057 | -34.1956 | -11.0369
6.3118 | -34.1959 | -11.0471
6.3179 | -34.1962 | -11.0572
6.324 | -34.1965 | -11.0674
6.3301 | -34.1968 | -11.0775
6.3362 | -34.1971 | -11.0875
6.3423 | -34.1975 | -11.0978
6.3484 | -34.1978 | -11.108
6.3545 | -34.1981 | -11.1179
6.3606 | -34.1984 | -11.1277
6.3667 | -34.1987 | -11.1377
6.3728 | -34.199 | -11.1479
6.379 | -34.1993 | -11.158
6.3851 | -34.1996 | -11.1679
6.3912 | -34.1999 | -11.1779
6.3973 | -34.2002 | -11.1881
6.4034 | -34.2005 | -11.198
6.4095 | -34.2008 | -11.2081
6.4156 | -34.2011 | -11.2182
6.4217 | -34.2014 | -11.2283
6.4278 | -34.2017 | -11.2384
6.4339 | -34.2021 | -11.2484
6.44 | -34.2024 | -11.2584
6.4461 | -34.2027 | -11.2687
6.4522 | -34.203 | -11.2785
6.4583 | -34.2032 | -11.2887
6.4644 | -34.2035 | -11.2989
6.4705 | -34.2039 | -11.3087
6.4766 | -34.2042 | -11.3188
6.4827 | -34.2045 | -11.3289
6.4888 | -34.2047 | -11.3389
6.4949 | -34.2051 | -11.3487
6.501 | -34.2054 | -11.3588
6.5071 | -34.2057 | -11.3687
6.5132 | -34.206 | -11.3788
6.5194 | -34.2063 | -11.389
6.5255 | -34.2066 | -11.3988
6.5316 | -34.2069 | -11.409
6.5377 | -34.2072 | -11.419
6.5438 | -34.2075 | -11.429
6.5499 | -34.2078 | -11.4389
6.556 | -34.2081 | -11.4489
6.5621 | -34.2084 | -11.4592
6.5682 | -34.2087 | -11.4693
6.5743 | -34.2091 | -11.4793
6.5804 | -34.2094 | -11.4895
6.5865 | -34.2097 | -11.4996
6.5926 | -34.21 | -11.5095
6.5987 | -34.2103 | -11.5193
6.6048 | -34.2106 | -11.5294
6.6109 | -34.211 | -11.5396
6.617 | -34.2112 | -11.5497
6.6231 | -34.2115 | -11.5598
6.6292 | -34.2119 | -11.5698
6.6353 | -34.2122 | -11.5802
6.6414 | -34.2125 | -11.5901
6.6475 | -34.2128 | -11.6002
6.6536 | -34.2131 | -11.6102
6.6597 | -34.2134 | -11.62
6.6659 | -34.2137 | -11.6301
6.672 | -34.214 | -11.64
6.6781 | -34.2143 | -11.6501
6.6842 | -34.2147 | -11.66
6.6903 | -34.215 | -11.67
6.6964 | -34.2153 | -11.68
6.7025 | -34.2156 | -11.6898
6.7086 | -34.2159 | -11.6999
6.7147 | -34.2162 | -11.71
6.7208 | -34.2165 | -11.7201
6.7269 | -34.2168 | -11.73
6.733 | -34.2171 | -11.7398
6.7391 | -34.2175 | -11.7499
6.7452 | -34.2178 | -11.7597
6.7513 | -34.2181 | -11.7695
6.7574 | -34.2184 | -11.7796
6.7635 | -34.2187 | -11.7894
6.7696 | -34.219 | -11.7993
6.7757 | -34.2194 | -11.8093
6.7818 | -34.2197 | -11.8193
6.7879 | -34.22 | -11.8292
6.794 | -34.2203 | -11.839
6.8001 | -34.2206 | -11.849
6.8063 | -34.2209 | -11.859
6.8124 | -34.2212 | -11.8692
6.8185 | -34.2215 | -11.8793
6.8246 | -34.2218 | -11.8892
6.8307 | -34.2221 | -11.8991
6.8368 | -34.2224 | -11.9093
6.8429 | -34.2228 | -11.9194
6.849 | -34.2231 | -11.9294
6.8551 | -34.2234 | -11.9393
6.8612 | -34.2237 | -11.9493
6.8673 | -34.224 | -11.9592
6.8734 | -34.2243 | -11.9694
6.8795 | -34.2247 | -11.9793
6.8856 | -34.225 | -11.9893
6.8917 | -34.2253 | -11.9993
6.8978 | -34.2256 | -12.009
6.9039 | -34.2259 | -12.0187
6.91 | -34.2263 | -12.0288
6.9161 | -34.2266 | -12.0385
6.9222 | -34.2269 | -12.0485
6.9283 | -34.2272 | -12.0582
6.9344 | -34.2276 | -12.0681
6.9405 | -34.2279 | -12.078
6.9466 | -34.2282 | -12.0882
6.9528 | -34.2286 | -12.0983
6.9589 | -34.2289 | -12.1084
6.965 | -34.2292 | -12.1185
6.9711 | -34.2296 | -12.1285
6.9772 | -34.23 | -12.1385
6.9833 | -34.2303 | -12.1484
6.9894 | -34.2306 | -12.1586
6.9955 | -34.2309 | -12.1683
7.0016 | -34.2312 | -12.1783
7.0077 | -34.2316 | -12.1882
7.0138 | -34.2319 | -12.1981
7.0199 | -34.2322 | -12.2079
7.026 | -34.2325 | -12.2178
7.0321 | -34.2329 | -12.2276
7.0382 | -34.2332 | -12.2372
7.0443 | -34.2335 | -12.2473
7.0504 | -34.2339 | -12.2573
7.0565 | -34.2342 | -12.2672
7.0626 | -34.2345 | -12.2769
7.0687 | -34.2348 | -12.2865
7.0748 | -34.2352 | -12.2961
7.0809 | -34.2355 | -12.306
7.087 | -34.2358 | -12.316
7.0932 | -34.2361 | -12.326
7.0993 | -34.2365 | -12.3359
7.1054 | -34.2368 | -12.3456
7.1115 | -34.2372 | -12.3553
7.1176 | -34.2375 | -12.3651
7.1237 | -34.2378 | -12.375
7.1298 | -34.2381 | -12.3848
7.1359 | -34.2384 | -12.3947
7.142 | -34.2388 | -12.4043
7.1481 | -34.2391 | -12.414
7.1542 | -34.2394 | -12.4239
7.1603 | -34.2397 | -12.4338
7.1664 | -34.2401 | -12.4436
7.1725 | -34.2404 | -12.4534
7.1786 | -34.2408 | -12.4633
7.1847 | -34.2411 | -12.4732
7.1908 | -34.2414 | -12.4832
7.1969 | -34.2417 | -12.4932
7.203 | -34.2421 | -12.5029
7.2091 | -34.2424 | -12.5126
7.2152 | -34.2427 | -12.5225
7.2213 | -34.2431 | -12.5321
7.2274 | -34.2434 | -12.5419
7.2335 | -34.2437 | -12.5517
7.2397 | -34.2441 | -12.5614
7.2458 | -34.2445 | -12.571
7.2519 | -34.2448 | -12.5808
7.258 | -34.2451 | -12.5908
7.2641 | -34.2455 | -12.6007
7.2702 | -34.2458 | -12.6107
7.2763 | -34.2461 | -12.6205
7.2824 | -34.2465 | -12.6304
7.2885 | -34.2468 | -12.6402
7.2946 | -34.2471 | -12.65
7.3007 | -34.2474 | -12.66
7.3068 | -34.2478 | -12.67
7.3129 | -34.2481 | -12.6798
7.319 | -34.2485 | -12.6896
7.3251 | -34.2488 | -12.6993
7.3312 | -34.2491 | -12.7091
7.3373 | -34.2495 | -12.7187
7.3434 | -34.2498 | -12.7284
7.3495 | -34.2502 | -12.738
7.3556 | -34.2505 | -12.7479
7.3617 | -34.2509 | -12.7579
7.3678 | -34.2512 | -12.7676
7.3739 | -34.2515 | -12.7776
7.3801 | -34.2519 | -12.7876
7.3862 | -34.2522 | -12.7974
7.3923 | -34.2526 | -12.8074
7.3984 | -34.2529 | -12.8169
7.4045 | -34.2533 | -12.8267
7.4106 | -34.2536 | -12.8364
7.4167 | -34.2539 | -12.8463
7.4228 | -34.2543 | -12.8562
7.4289 | -34.2547 | -12.866
7.435 | -34.255 | -12.8755
7.4411 | -34.2553 | -12.8854
7.4472 | -34.2557 | -12.8952
7.4533 | -34.256 | -12.9047
7.4594 | -34.2564 | -12.9147
7.4655 | -34.2567 | -12.9244
7.4716 | -34.2571 | -12.9344
7.4777 | -34.2574 | -12.9441
7.4838 | -34.2577 | -12.9537
7.4899 | -34.2581 | -12.9633
7.496 | -34.2584 | -12.9729
7.5021 | -34.2588 | -12.9826
7.5082 | -34.2591 | -12.9924
7.5143 | -34.2594 | -13.0023
7.5204 | -34.2598 | -13.0121
7.5266 | -34.2601 | -13.0217
7.5327 | -34.2604 | -13.0314
7.5388 | -34.2608 | -13.0414
7.5449 | -34.2611 | -13.0514
7.551 | -34.2615 | -13.061
7.5571 | -34.2618 | -13.0709
7.5632 | -34.2622 | -13.0807
7.5693 | -34.2626 | -13.0903
7.5754 | -34.2629 | -13.0999
7.5815 | -34.2632 | -13.1097
7.5876 | -34.2636 | -13.1197
7.5937 | -34.2639 | -13.1296
7.5998 | -34.2643 | -13.1393
7.6059 | -34.2647 | -13.1493
7.612 | -34.265 | -13.159
7.6181 | -34.2654 | -13.1687
7.6242 | -34.2657 | -13.1787
7.6303 | -34.2661 | -13.1886
7.6364 | -34.2664 | -13.1985
7.6425 | -34.2667 | -13.2083
7.6486 | -34.2671 | -13.2181
7.6547 | -34.2674 | -13.2278
7.6608 | -34.2678 | -13.2375
7.667 | -34.2681 | -13.2474
7.6731 | -34.2684 | -13.2572
7.6792 | -34.2688 | -13.2672
7.6853 | -34.2691 | -13.2771
7.6914 | -34.2695 | -13.287
7.6975 | -34.2699 | -13.2968
7.7036 | -34.2702 | -13.3066
7.7097 | -34.2706 | -13.3166
7.7158 | -34.2709 | -13.3264
7.7219 | -34.2713 | -13.336
7.728 | -34.2716 | -13.3459
7.7341 | -34.272 | -13.3557
7.7402 | -34.2723 | -13.3656
7.7463 | -34.2727 | -13.3754
7.7524 | -34.273 | -13.3855
7.7585 | -34.2734 | -13.3953
7.7646 | -34.2737 | -13.4051
7.7707 | -34.2741 | -13.415
7.7768 | -34.2744 | -13.4249
7.7829 | -34.2747 | -13.435
7.789 | -34.2751 | -13.445
7.7951 | -34.2754 | -13.4548
7.8012 | -34.2758 | -13.4646
7.8073 | -34.2762 | -13.4746
7.8135 | -34.2765 | -13.4844
7.8196 | -34.2769 | -13.4944
7.8257 | -34.2772 | -13.5044
7.8318 | -34.2775 | -13.5144
7.8379 | -34.2779 | -13.5241
7.844 | -34.2783 | -13.5338
7.8501 | -34.2786 | -13.5439
7.8562 | -34.2789 | -13.5539
7.8623 | -34.2793 | -13.5637
7.8684 | -34.2797 | -13.5736
7.8745 | -34.28 | -13.5835
7.8806 | -34.2803 | -13.5935
7.8867 | -34.2807 | -13.6031
7.8928 | -34.2811 | -13.6127
7.8989 | -34.2814 | -13.6226
7.905 | -34.2818 | -13.6325
7.9111 | -34.2821 | -13.6424
7.9172 | -34.2824 | -13.6523
7.9233 | -34.2828 | -13.6622
7.9294 | -34.2831 | -13.6722
7.9355 | -34.2835 | -13.6818
7.9416 | -34.2838 | -13.6917
7.9477 | -34.2841 | -13.7014
7.9539 | -34.2845 | -13.7112
7.96 | -34.2848 | -13.7211
7.9661 | -34.2852 | -13.7309
7.9722 | -34.2856 | -13.7406
7.9783 | -34.2859 | -13.7507
7.9844 | -34.2863 | -13.7604
7.9905 | -34.2866 | -13.7705
7.9966 | -34.2869 | -13.7804
8.0027 | -34.2873 | -13.7903
8.0088 | -34.2876 | -13.8001
8.0149 | -34.2879 | -13.81
8.021 | -34.2883 | -13.8199
8.0271 | -34.2887 | -13.83
8.0332 | -34.289 | -13.8397
8.0393 | -34.2894 | -13.8497
8.0454 | -34.2897 | -13.8598
8.0515 | -34.2901 | -13.8696
8.0576 | -34.2904 | -13.8797
8.0637 | -34.2908 | -13.8894
8.0698 | -34.2911 | -13.8992
8.0759 | -34.2914 | -13.9089
8.082 | -34.2918 | -13.9187
8.0881 | -34.2921 | -13.9286
8.0942 | -34.2925 | -13.9385
8.1004 | -34.2929 | -13.9485
8.1065 | -34.2932 | -13.9582
8.1126 | -34.2936 | -13.9681
8.1187 | -34.2939 | -13.9781
8.1248 | -34.2942 | -13.9877
8.1309 | -34.2946 | -13.9975
8.137 | -34.295 | -14.0071
8.1431 | -34.2953 | -14.0172
8.1492 | -34.2957 | -14.027
8.1553 | -34.296 | -14.0369
8.1614 | -34.2963 | -14.047
8.1675 | -34.2967 | -14.0568
8.1736 | -34.297 | -14.0668
8.1797 | -34.2974 | -14.0767
8.1858 | -34.2977 | -14.0866
8.1919 | -34.2981 | -14.0965
8.198 | -34.2984 | -14.1062
8.2041 | -34.2988 | -14.1161
8.2102 | -34.2991 | -14.126
8.2163 | -34.2995 | -14.1357
8.2224 | -34.2998 | -14.1455
8.2285 | -34.3001 | -14.1551
8.2346 | -34.3005 | -14.1649
8.2408 | -34.3009 | -14.1747
8.2469 | -34.3012 | -14.1846
8.253 | -34.3016 | -14.1943
8.2591 | -34.3019 | -14.2043
8.2652 | -34.3023 | -14.2142
8.2713 | -34.3027 | -14.2241
8.2774 | -34.303 | -14.234
8.2835 | -34.3034 | -14.2441
8.2896 | -34.3037 | -14.2537
8.2957 | -34.3041 | -14.2637
8.3018 | -34.3044 | -14.2736
8.3079 | -34.3048 | -14.2833
8.314 | -34.3051 | -14.2932
8.3201 | -34.3055 | -14.303
8.3262 | -34.3058 | -14.3129
8.3323 | -34.3062 | -14.3226
8.3384 | -34.3066 | -14.3326
8.3445 | -34.3069 | -14.3423
8.3506 | -34.3072 | -14.3522
8.3567 | -34.3076 | -14.3621
8.3628 | -34.3079 | -14.372
8.3689 | -34.3083 | -14.3817
8.375 | -34.3086 | -14.3916
8.3812 | -34.309 | -14.4012
8.3873 | -34.3093 | -14.4111
8.3934 | -34.3097 | -14.421
8.3995 | -34.31 | -14.4307
8.4056 | -34.3104 | -14.4405
8.4117 | -34.3107 | -14.4504
8.4178 | -34.3111 | -14.4603
8.4239 | -34.3114 | -14.47
8.43 | -34.3118 | -14.4798
8.4361 | -34.3121 | -14.4898
8.4422 | -34.3125 | -14.4995
8.4483 | -34.3128 | -14.5092
8.4544 | -34.3132 | -14.5191
8.4605 | -34.3135 | -14.5287
8.4666 | -34.3138 | -14.5387
8.4727 | -34.3142 | -14.5486
8.4788 | -34.3145 | -14.5584
8.4849 | -34.3149 | -14.5682
8.491 | -34.3152 | -14.578
8.4971 | -34.3156 | -14.5876
8.5032 | -34.3159 | -14.5974
8.5093 | -34.3163 | -14.607
8.5154 | -34.3167 | -14.6169
8.5215 | -34.317 | -14.6265
8.5277 | -34.3174 | -14.6364
8.5338 | -34.3177 | -14.6461
8.5399 | -34.3181 | -14.6559
8.546 | -34.3185 | -14.6656
8.5521 | -34.3188 | -14.6753
8.5582 | -34.3191 | -14.6851
8.5643 | -34.3195 | -14.6947
8.5704 | -34.3198 | -14.7043
8.5765 | -34.3202 | -14.7142
8.5826 | -34.3206 | -14.7236
8.5887 | -34.3209 | -14.7333
8.5948 | -34.3213 | -14.7431
8.6009 | -34.3217 | -14.7527
8.607 | -34.322 | -14.7625
8.6131 | -34.3224 | -14.7723
8.6192 | -34.3228 | -14.782
8.6253 | -34.3231 | -14.7916
8.6314 | -34.3235 | -14.8013
8.6375 | -34.3238 | -14.8111
8.6436 | -34.3242 | -14.8209
8.6497 | -34.3245 | -14.8306
8.6558 | -34.3249 | -14.8403
8.6619 | -34.3252 | -14.8501
8.6681 | -34.3256 | -14.86
8.6742 | -34.326 | -14.8697
8.6803 | -34.3263 | -14.8793
8.6864 | -34.3267 | -14.8888
8.6925 | -34.327 | -14.8983
8.6986 | -34.3274 | -14.9081
8.7047 | -34.3277 | -14.918
8.7108 | -34.3281 | -14.9277
8.7169 | -34.3285 | -14.9375
8.723 | -34.3288 | -14.9471
8.7291 | -34.3292 | -14.9566
8.7352 | -34.3295 | -14.9665
8.7413 | -34.3299 | -14.9763
8.7474 | -34.3302 | -14.9858
8.7535 | -34.3306 | -14.9956
8.7596 | -34.331 | -15.0053
8.7657 | -34.3313 | -15.0151
8.7718 | -34.3316 | -15.0249
8.7779 | -34.332 | -15.0347
8.784 | -34.3324 | -15.0445
8.7901 | -34.3328 | -15.0542
8.7962 | -34.3331 | -15.0637
8.8023 | -34.3335 | -15.0733
8.8084 | -34.3338 | -15.083
8.8146 | -34.3342 | -15.0927
8.8207 | -34.3346 | -15.1023
8.8268 | -34.335 | -15.1117
8.8329 | -34.3353 | -15.1214
8.839 | -34.3357 | -15.1312
8.8451 | -34.336 | -15.1407
8.8512 | -34.3364 | -15.1503
8.8573 | -34.3367 | -15.16
8.8634 | -34.3371 | -15.1695
8.8695 | -34.3375 | -15.1794
8.8756 | -34.3378 | -15.1891
8.8817 | -34.3382 | -15.1987
8.8878 | -34.3386 | -15.2085
8.8939 | -34.3389 | -15.2182
8.9 | -34.3393 | -15.2278
8.9061 | -34.3397 | -15.2374
8.9122 | -34.34 | -15.2471
8.9183 | -34.3404 | -15.2568
8.9244 | -34.3408 | -15.2666
8.9305 | -34.3412 | -15.276
8.9366 | -34.3416 | -15.2857
8.9427 | -34.3419 | -15.2954
8.9488 | -34.3423 | -15.305
8.955 | -34.3427 | -15.3146
8.9611 | -34.343 | -15.3244
8.9672 | -34.3435 | -15.3341
8.9733 | -34.3438 | -15.3435
8.9794 | -34.3442 | -15.3531
8.9855 | -34.3446 | -15.3628
8.9916 | -34.345 | -15.3726
8.9977 | -34.3453 | -15.3822
9.0038 | -34.3457 | -15.3918
9.0099 | -34.3461 | -15.4014
9.016 | -34.3465 | -15.4111
9.0221 | -34.3468 | -15.4206
9.0282 | -34.3472 | -15.43
9.0343 | -34.3476 | -15.4396
9.0404 | -34.348 | -15.4493
9.0465 | -34.3484 | -15.4591
9.0526 | -34.3488 | -15.4686
9.0587 | -34.3492 | -15.4782
9.0648 | -34.3496 | -15.4877
9.0709 | -34.3499 | -15.4973
9.077 | -34.3504 | -15.507
9.0831 | -34.3507 | -15.5167
9.0892 | -34.3511 | -15.5264
9.0953 | -34.3515 | -15.5361
9.1015 | -34.3519 | -15.5457
9.1076 | -34.3523 | -15.5552
9.1137 | -34.3526 | -15.5648
9.1198 | -34.3531 | -15.5744
9.1259 | -34.3534 | -15.5839
9.132 | -34.3538 | -15.5937
9.1381 | -34.3542 | -15.6032
9.1442 | -34.3546 | -15.6125
9.1503 | -34.355 | -15.6223
9.1564 | -34.3554 | -15.6318
9.1625 | -34.3558 | -15.6415
9.1686 | -34.3562 | -15.6511
9.1747 | -34.3566 | -15.6607
9.1808 | -34.357 | -15.6702
9.1869 | -34.3574 | -15.6799
9.193 | -34.3578 | -15.6895
9.1991 | -34.3581 | -15.6991
9.2052 | -34.3585 | -15.7087
9.2113 | -34.3589 | -15.7183
9.2174 | -34.3594 | -15.728
9.2235 | -34.3598 | -15.7375
9.2296 | -34.3602 | -15.7471
9.2357 | -34.3605 | -15.7569
9.2419 | -34.3609 | -15.7664
9.248 | -34.3614 | -15.776
9.2541 | -34.3618 | -15.7856
9.2602 | -34.3622 | -15.7953
9.2663 | -34.3626 | -15.8047
9.2724 | -34.363 | -15.8142
9.2785 | -34.3634 | -15.8238
9.2846 | -34.3638 | -15.8333
9.2907 | -34.3642 | -15.8428
9.2968 | -34.3646 | -15.8526
9.3029 | -34.365 | -15.8623
9.309 | -34.3654 | -15.8719
9.3151 | -34.3657 | -15.8813
9.3212 | -34.3662 | -15.8907
9.3273 | -34.3666 | -15.9003
9.3334 | -34.367 | -15.9097
9.3395 | -34.3674 | -15.9192
9.3456 | -34.3677 | -15.9286
9.3517 | -34.3681 | -15.9383
9.3578 | -34.3686 | -15.9478
9.3639 | -34.369 | -15.9572
9.37 | -34.3694 | -15.9668
9.3761 | -34.3698 | -15.9765
9.3822 | -34.3702 | -15.9861
9.3884 | -34.3706 | -15.9955
9.3945 | -34.371 | -16.0051
9.4006 | -34.3714 | -16.0146
9.4067 | -34.3718 | -16.0242
9.4128 | -34.3722 | -16.0338
9.4189 | -34.3727 | -16.0436
9.425 | -34.3731 | -16.053
9.4311 | -34.3735 | -16.0625
9.4372 | -34.3739 | -16.0721
9.4433 | -34.3743 | -16.0816
9.4494 | -34.3747 | -16.091
9.4555 | -34.3752 | -16.1006
9.4616 | -34.3756 | -16.1101
9.4677 | -34.376 | -16.1197
9.4738 | -34.3764 | -16.1292
9.4799 | -34.3768 | -16.1388
9.486 | -34.3772 | -16.1484
9.4921 | -34.3776 | -16.1578
9.4982 | -34.378 | -16.1672
9.5043 | -34.3784 | -16.1767
9.5104 | -34.3788 | -16.1861
9.5165 | -34.3793 | -16.1957
9.5226 | -34.3796 | -16.2054
9.5288 | -34.3801 | -16.2148
9.5349 | -34.3805 | -16.2244
9.541 | -34.3809 | -16.2337
9.5471 | -34.3813 | -16.2432
9.5532 | -34.3817 | -16.2528
9.5593 | -34.3821 | -16.2624
9.5654 | -34.3825 | -16.2719
9.5715 | -34.3829 | -16.2814
9.5776 | -34.3833 | -16.2911
9.5837 | -34.3837 | -16.3005
9.5898 | -34.3841 | -16.3101
9.5959 | -34.3845 | -16.3197
9.602 | -34.3849 | -16.329
9.6081 | -34.3854 | -16.3385
9.6142 | -34.3858 | -16.348
9.6203 | -34.3862 | -16.3575
9.6264 | -34.3866 | -16.3669
9.6325 | -34.387 | -16.3766
9.6386 | -34.3874 | -16.386
9.6447 | -34.3878 | -16.3955
9.6508 | -34.3882 | -16.4049
9.6569 | -34.3886 | -16.4145
9.663 | -34.389 | -16.4241
9.6691 | -34.3895 | -16.4336
9.6753 | -34.3899 | -16.443
9.6814 | -34.3903 | -16.4524
9.6875 | -34.3907 | -16.4617
9.6936 | -34.3911 | -16.471
9.6997 | -34.3915 | -16.4806
9.7058 | -34.392 | -16.49
9.7119 | -34.3924 | -16.4996
9.718 | -34.3928 | -16.5092
9.7241 | -34.3933 | -16.5187
9.7302 | -34.3936 | -16.5281
9.7363 | -34.3941 | -16.5375
9.7424 | -34.3945 | -16.5469
9.7485 | -34.3949 | -16.5564
9.7546 | -34.3953 | -16.5657
9.7607 | -34.3957 | -16.5754
9.7668 | -34.3961 | -16.585
9.7729 | -34.3966 | -16.5944
9.779 | -34.397 | -16.6037
9.7851 | -34.3974 | -16.6133
9.7912 | -34.3979 | -16.6227
9.7973 | -34.3983 | -16.6323
9.8034 | -34.3987 | -16.6415
9.8095 | -34.3991 | -16.6512
9.8157 | -34.3996 | -16.6606
9.8218 | -34.4 | -16.6703
9.8279 | -34.4004 | -16.6796
9.834 | -34.4008 | -16.6892
9.8401 | -34.4012 | -16.6986
9.8462 | -34.4016 | -16.7078
9.8523 | -34.402 | -16.7173
9.8584 | -34.4025 | -16.7267
9.8645 | -34.4029 | -16.7363
9.8706 | -34.4033 | -16.7459
9.8767 | -34.4036 | -16.7553
9.8828 | -34.4041 | -16.7646
9.8889 | -34.4045 | -16.7739
9.895 | -34.4049 | -16.7834
9.9011 | -34.4053 | -16.7929
9.9072 | -34.4057 | -16.8023
9.9133 | -34.4061 | -16.8117
9.9194 | -34.4065 | -16.821
9.9255 | -34.4069 | -16.8304
9.9316 | -34.4074 | -16.8398
9.9377 | -34.4078 | -16.8492
9.9438 | -34.4082 | -16.8587
9.9499 | -34.4086 | -16.8683
9.956 | -34.409 | -16.8777
9.9622 | -34.4094 | -16.8873
9.9683 | -34.4098 | -16.8967
9.9744 | -34.4102 | -16.9063
9.9805 | -34.4107 | -16.9157
9.9866 | -34.4111 | -16.9251
9.9927 | -34.4115 | -16.9344
9.9988 | -34.4119 | -16.944
10.0049 | -34.4123 | -16.9534
10.011 | -34.4127 | -16.9629
10.0171 | -34.4132 | -16.9722
10.0232 | -34.4136 | -16.9815
10.0293 | -34.414 | -16.991
10.0354 | -34.4145 | -17.0005
10.0415 | -34.4149 | -17.0099
10.0476 | -34.4153 | -17.0195
10.0537 | -34.4157 | -17.029
10.0598 | -34.4161 | -17.0383
10.0659 | -34.4165 | -17.0478
10.072 | -34.4169 | -17.0571
10.0781 | -34.4173 | -17.0664
10.0842 | -34.4178 | -17.0759
10.0903 | -34.4182 | -17.0853
10.0964 | -34.4186 | -17.0944
10.1026 | -34.419 | -17.1039
10.1087 | -34.4194 | -17.1132
10.1148 | -34.4198 | -17.1227
10.1209 | -34.4202 | -17.132
10.127 | -34.4206 | -17.1413
10.1331 | -34.4211 | -17.1506
10.1392 | -34.4214 | -17.1599
10.1453 | -34.4219 | -17.1692
10.1514 | -34.4223 | -17.1786
10.1575 | -34.4227 | -17.1881
10.1636 | -34.4231 | -17.1973
10.1697 | -34.4235 | -17.2068
10.1758 | -34.4239 | -17.216
10.1819 | -34.4244 | -17.2253
10.188 | -34.4248 | -17.2347
10.1941 | -34.4252 | -17.2441
10.2002 | -34.4256 | -17.2535
10.2063 | -34.426 | -17.2628
10.2124 | -34.4264 | -17.2723
10.2185 | -34.4268 | -17.2817
10.2246 | -34.4272 | -17.291
10.2307 | -34.4276 | -17.3004
10.2368 | -34.4281 | -17.3095
10.2429 | -34.4284 | -17.3189
10.2491 | -34.4289 | -17.3283
10.2552 | -34.4293 | -17.3377
10.2613 | -34.4297 | -17.3471
10.2674 | -34.4301 | -17.3565
10.2735 | -34.4305 | -17.3661
10.2796 | -34.4309 | -17.3754
10.2857 | -34.4314 | -17.3849
10.2918 | -34.4318 | -17.3944
10.2979 | -34.4322 | -17.4038
10.304 | -34.4326 | -17.4129
10.3101 | -34.433 | -17.4223
10.3162 | -34.4334 | -17.4318
10.3223 | -34.4339 | -17.441
10.3284 | -34.4343 | -17.4504
10.3345 | -34.4347 | -17.4596
10.3406 | -34.4351 | -17.4689
10.3467 | -34.4355 | -17.4782
10.3528 | -34.4359 | -17.4875
10.3589 | -34.4363 | -17.4968
10.365 | -34.4367 | -17.5063
10.3711 | -34.4371 | -17.5157
10.3772 | -34.4375 | -17.5251
10.3833 | -34.438 | -17.5344
10.3895 | -34.4384 | -17.5438
10.3956 | -34.4388 | -17.5531
10.4017 | -34.4392 | -17.5623
10.4078 | -34.4396 | -17.5717
10.4139 | -34.4401 | -17.581
10.42 | -34.4405 | -17.5904
10.4261 | -34.4409 | -17.5997
10.4322 | -34.4413 | -17.609
10.4383 | -34.4417 | -17.6184
10.4444 | -34.4421 | -17.6275
10.4505 | -34.4425 | -17.6369
10.4566 | -34.4429 | -17.6462
10.4627 | -34.4433 | -17.6554
10.4688 | -34.4437 | -17.6646
10.4749 | -34.4441 | -17.674
10.481 | -34.4445 | -17.6833
10.4871 | -34.445 | -17.6925
10.4932 | -34.4454 | -17.7017
10.4993 | -34.4458 | -17.7111
10.5054 | -34.4462 | -17.7204
10.5115 | -34.4466 | -17.7297
10.5176 | -34.4471 | -17.7388
10.5237 | -34.4475 | -17.7483
10.5298 | -34.4479 | -17.7575
10.536 | -34.4483 | -17.7669
10.5421 | -34.4487 | -17.776
10.5482 | -34.4491 | -17.7854
10.5543 | -34.4496 | -17.7946
10.5604 | -34.45 | -17.8037
10.5665 | -34.4504 | -17.8129
10.5726 | -34.4508 | -17.8221
10.5787 | -34.4512 | -17.8314
10.5848 | -34.4516 | -17.8404
10.5909 | -34.452 | -17.8498
10.597 | -34.4524 | -17.8591
10.6031 | -34.4529 | -17.8685
10.6092 | -34.4533 | -17.8778
10.6153 | -34.4537 | -17.8872
10.6214 | -34.4541 | -17.8962
10.6275 | -34.4545 | -17.9054
10.6336 | -34.4549 | -17.9148
10.6397 | -34.4553 | -17.9241
10.6458 | -34.4557 | -17.9333
10.6519 | -34.4561 | -17.9427
10.658 | -34.4566 | -17.9519
10.6641 | -34.457 | -17.9613
10.6702 | -34.4574 | -17.9706
10.6764 | -34.4578 | -17.9799
10.6825 | -34.4582 | -17.9891
10.6886 | -34.4586 | -17.9982
10.6947 | -34.4591 | -18.0074
10.7008 | -34.4595 | -18.0167
10.7069 | -34.4599 | -18.026
10.713 | -34.4603 | -18.0353
10.7191 | -34.4608 | -18.0445
10.7252 | -34.4612 | -18.0535
10.7313 | -34.4616 | -18.0627
10.7374 | -34.462 | -18.072
10.7435 | -34.4624 | -18.0812
10.7496 | -34.4628 | -18.0903
10.7557 | -34.4632 | -18.0996
10.7618 | -34.4637 | -18.1086
10.7679 | -34.4641 | -18.1179
10.774 | -34.4645 | -18.1272
10.7801 | -34.4649 | -18.1364
10.7862 | -34.4653 | -18.1455
10.7923 | -34.4657 | -18.1548
10.7984 | -34.4661 | -18.1638
10.8045 | -34.4666 | -18.1729
10.8106 | -34.467 | -18.1822
10.8168 | -34.4674 | -18.1913
10.8229 | -34.4678 | -18.2005
10.829 | -34.4682 | -18.2097
10.8351 | -34.4686 | -18.2186
10.8412 | -34.469 | -18.2277
10.8473 | -34.4695 | -18.2367
10.8534 | -34.4699 | -18.2459
10.8595 | -34.4703 | -18.2551
10.8656 | -34.4707 | -18.2641
10.8717 | -34.4711 | -18.2732
10.8778 | -34.4716 | -18.2824
10.8839 | -34.472 | -18.2916
10.89 | -34.4724 | -18.3008
10.8961 | -34.4728 | -18.3099
10.9022 | -34.4732 | -18.319
10.9083 | -34.4737 | -18.3281
10.9144 | -34.4741 | -18.3372
10.9205 | -34.4745 | -18.3464
10.9266 | -34.4749 | -18.3555
10.9327 | -34.4753 | -18.3646
10.9388 | -34.4758 | -18.3735
10.9449 | -34.4762 | -18.3826
10.951 | -34.4766 | -18.3918
10.9571 | -34.477 | -18.401
10.9633 | -34.4774 | -18.4102
10.9694 | -34.4778 | -18.4196
10.9755 | -34.4783 | -18.4288
10.9816 | -34.4787 | -18.438
10.9877 | -34.4791 | -18.447
10.9938 | -34.4795 | -18.4564
10.9999 | -34.48 | -18.4654
11.006 | -34.4804 | -18.4742
11.0121 | -34.4808 | -18.4834
11.0182 | -34.4812 | -18.4925
11.0243 | -34.4816 | -18.5015
11.0304 | -34.4821 | -18.5107
11.0365 | -34.4825 | -18.5197
11.0426 | -34.4829 | -18.5286
11.0487 | -34.4834 | -18.5378
11.0548 | -34.4838 | -18.5468
11.0609 | -34.4842 | -18.556
11.067 | -34.4847 | -18.5651
11.0731 | -34.4851 | -18.5741
11.0792 | -34.4855 | -18.5831
11.0853 | -34.486 | -18.5922
11.0914 | -34.4864 | -18.6014
11.0975 | -34.4868 | -18.6105
11.1037 | -34.4872 | -18.6195
11.1098 | -34.4876 | -18.6288
11.1159 | -34.4881 | -18.6376
11.122 | -34.4885 | -18.6466
11.1281 | -34.4889 | -18.6557
11.1342 | -34.4894 | -18.6648
11.1403 | -34.4898 | -18.674
11.1464 | -34.4902 | -18.6831
11.1525 | -34.4906 | -18.6919
11.1586 | -34.491 | -18.7011
11.1647 | -34.4915 | -18.7102
11.1708 | -34.4919 | -18.7192
11.1769 | -34.4923 | -18.7283
11.183 | -34.4928 | -18.7372
11.1891 | -34.4932 | -18.7463
11.1952 | -34.4936 | -18.7553
11.2013 | -34.4941 | -18.7644
11.2074 | -34.4945 | -18.7735
11.2135 | -34.4949 | -18.7826
11.2196 | -34.4954 | -18.7915
11.2257 | -34.4958 | -18.8006
11.2318 | -34.4962 | -18.8096
11.2379 | -34.4967 | -18.8187
11.244 | -34.4971 | -18.8279
11.2502 | -34.4976 | -18.8368
11.2563 | -34.498 | -18.8459
11.2624 | -34.4985 | -18.8548
11.2685 | -34.4989 | -18.8637
11.2746 | -34.4993 | -18.8727
11.2807 | -34.4997 | -18.8818
11.2868 | -34.5002 | -18.891
11.2929 | -34.5006 | -18.8997
11.299 | -34.5011 | -18.9088
11.3051 | -34.5015 | -18.918
11.3112 | -34.5019 | -18.9271
11.3173 | -34.5024 | -18.9359
11.3234 | -34.5028 | -18.9447
11.3295 | -34.5033 | -18.9539
11.3356 | -34.5037 | -18.963
11.3417 | -34.5041 | -18.972
11.3478 | -34.5045 | -18.9811
11.3539 | -34.5049 | -18.9903
11.36 | -34.5054 | -18.999
11.3661 | -34.5058 | -19.008
11.3722 | -34.5062 | -19.0167
11.3783 | -34.5067 | -19.0258
11.3844 | -34.5071 | -19.0348
11.3906 | -34.5075 | -19.0438
11.3967 | -34.508 | -19.0526
11.4028 | -34.5084 | -19.0618
11.4089 | -34.5088 | -19.0706
11.415 | -34.5092 | -19.0796
11.4211 | -34.5097 | -19.0887
11.4272 | -34.5101 | -19.0977
11.4333 | -34.5105 | -19.1064
11.4394 | -34.5109 | -19.1153
11.4455 | -34.5113 | -19.1243
11.4516 | -34.5118 | -19.1333
11.4577 | -34.5122 | -19.1423
11.4638 | -34.5127 | -19.1511
11.4699 | -34.5131 | -19.1601
11.476 | -34.5135 | -19.1692
11.4821 | -34.514 | -19.1782
11.4882 | -34.5144 | -19.1872
11.4943 | -34.5148 | -19.1963
11.5004 | -34.5153 | -19.2051
11.5065 | -34.5157 | -19.214
11.5126 | -34.5161 | -19.2228
11.5187 | -34.5165 | -19.2317
11.5248 | -34.517 | -19.2408
11.5309 | -34.5174 | -19.2497
11.5371 | -34.5179 | -19.2585
11.5432 | -34.5183 | -19.2676
11.5493 | -34.5187 | -19.2767
11.5554 | -34.5191 | -19.2857
11.5615 | -34.5196 | -19.2948
11.5676 | -34.52 | -19.3038
11.5737 | -34.5204 | -19.3125
11.5798 | -34.5209 | -19.3213
11.5859 | -34.5213 | -19.3303
11.592 | -34.5218 | -19.3392
11.5981 | -34.5222 | -19.348
11.6042 | -34.5226 | -19.3568
11.6103 | -34.523 | -19.3656
11.6164 | -34.5235 | -19.3745
11.6225 | -34.5239 | -19.3834
11.6286 | -34.5243 | -19.3923
11.6347 | -34.5248 | -19.4013
11.6408 | -34.5252 | -19.4101
11.6469 | -34.5256 | -19.419
11.653 | -34.5261 | -19.4278
11.6591 | -34.5265 | -19.4368
11.6652 | -34.5269 | -19.4456
11.6713 | -34.5274 | -19.4544
11.6775 | -34.5279 | -19.4631
11.6836 | -34.5283 | -19.472
11.6897 | -34.5288 | -19.4809
11.6958 | -34.5292 | -19.4898
11.7019 | -34.5296 | -19.4989
11.708 | -34.5301 | -19.5077
11.7141 | -34.5305 | -19.5165
11.7202 | -34.531 | -19.5253
11.7263 | -34.5314 | -19.5342
11.7324 | -34.5318 | -19.5432
11.7385 | -34.5322 | -19.5521
11.7446 | -34.5326 | -19.5609
11.7507 | -34.5331 | -19.5695
11.7568 | -34.5335 | -19.5784
11.7629 | -34.534 | -19.5872
11.769 | -34.5344 | -19.596
11.7751 | -34.5348 | -19.6048
11.7812 | -34.5353 | -19.6134
11.7873 | -34.5357 | -19.6223
11.7934 | -34.5362 | -19.6312
11.7995 | -34.5366 | -19.64
11.8056 | -34.5371 | -19.6487
11.8117 | -34.5375 | -19.6576
11.8178 | -34.5379 | -19.6662
11.824 | -34.5384 | -19.675
11.8301 | -34.5388 | -19.6838
11.8362 | -34.5392 | -19.6925
11.8423 | -34.5397 | -19.7013
11.8484 | -34.5401 | -19.7102
11.8545 | -34.5406 | -19.7189
11.8606 | -34.541 | -19.7276
11.8667 | -34.5415 | -19.7365
11.8728 | -34.5419 | -19.7452
11.8789 | -34.5423 | -19.7539
11.885 | -34.5428 | -19.7628
11.8911 | -34.5432 | -19.7714
11.8972 | -34.5437 | -19.7802
11.9033 | -34.5441 | -19.7889
11.9094 | -34.5445 | -19.7978
11.9155 | -34.545 | -19.8066
11.9216 | -34.5454 | -19.8152
11.9277 | -34.5458 | -19.824
11.9338 | -34.5463 | -19.8328
11.9399 | -34.5467 | -19.8417
11.946 | -34.5472 | -19.8504
11.9521 | -34.5476 | -19.8592
11.9582 | -34.548 | -19.8678
11.9644 | -34.5485 | -19.8765
11.9705 | -34.5489 | -19.8852
11.9766 | -34.5494 | -19.894
11.9827 | -34.5499 | -19.9027
11.9888 | -34.5503 | -19.9115
11.9949 | -34.5508 | -19.9199
12.001 | -34.5512 | -19.9287
12.0071 | -34.5517 | -19.9374
12.0132 | -34.5521 | -19.9461
12.0193 | -34.5525 | -19.9549
12.0254 | -34.553 | -19.9637
12.0315 | -34.5534 | -19.9722
12.0376 | -34.5539 | -19.9808
12.0437 | -34.5543 | -19.9896
12.0498 | -34.5547 | -19.9984
12.0559 | -34.5552 | -20.0071
12.062 | -34.5557 | -20.0155
12.0681 | -34.5561 | -20.0243
12.0742 | -34.5566 | -20.0328
12.0803 | -34.557 | -20.0417
12.0864 | -34.5574 | -20.0505
12.0925 | -34.5579 | -20.0593
12.0986 | -34.5583 | -20.0679
12.1047 | -34.5588 | -20.0765
12.1109 | -34.5592 | -20.085
12.117 | -34.5597 | -20.0936
12.1231 | -34.5601 | -20.1022
12.1292 | -34.5606 | -20.1108
12.1353 | -34.561 | -20.1194
12.1414 | -34.5615 | -20.1283
12.1475 | -34.5619 | -20.137
12.1536 | -34.5623 | -20.1458
12.1597 | -34.5628 | -20.1544
12.1658 | -34.5633 | -20.1631
12.1719 | -34.5637 | -20.1715
12.178 | -34.5641 | -20.1802
12.1841 | -34.5646 | -20.189
12.1902 | -34.565 | -20.1977
12.1963 | -34.5655 | -20.2065
12.2024 | -34.566 | -20.2149
12.2085 | -34.5664 | -20.2236
12.2146 | -34.5669 | -20.2323
12.2207 | -34.5673 | -20.241
12.2268 | -34.5678 | -20.2496
12.2329 | -34.5682 | -20.2582
12.239 | -34.5687 | -20.2666
12.2451 | -34.5692 | -20.2753
12.2513 | -34.5696 | -20.284
12.2574 | -34.57 | -20.2927
12.2635 | -34.5705 | -20.3014
12.2696 | -34.5709 | -20.3099
12.2757 | -34.5714 | -20.3183
12.2818 | -34.5719 | -20.3269
12.2879 | -34.5723 | -20.3355
12.294 | -34.5728 | -20.3441
12.3001 | -34.5732 | -20.3526
12.3062 | -34.5737 | -20.3612
12.3123 | -34.5741 | -20.3696
12.3184 | -34.5746 | -20.3781
12.3245 | -34.575 | -20.3866
12.3306 | -34.5755 | -20.3953
12.3367 | -34.5759 | -20.4039
12.3428 | -34.5763 | -20.4125
12.3489 | -34.5768 | -20.4209
12.355 | -34.5773 | -20.4297
12.3611 | -34.5777 | -20.4385
12.3672 | -34.5781 | -20.4469
12.3733 | -34.5786 | -20.4556
12.3794 | -34.5791 | -20.464
12.3855 | -34.5795 | -20.4725
12.3916 | -34.5799 | -20.481
12.3978 | -34.5804 | -20.4896
12.4039 | -34.5809 | -20.4983
12.41 | -34.5813 | -20.5069
12.4161 | -34.5818 | -20.5151
12.4222 | -34.5823 | -20.5237
12.4283 | -34.5827 | -20.5322
12.4344 | -34.5832 | -20.5409
12.4405 | -34.5836 | -20.5495
12.4466 | -34.5841 | -20.558
12.4527 | -34.5845 | -20.5665
12.4588 | -34.585 | -20.5751
12.4649 | -34.5854 | -20.5838
12.471 | -34.5859 | -20.5925
12.4771 | -34.5863 | -20.601
12.4832 | -34.5867 | -20.6095
12.4893 | -34.5872 | -20.6179
12.4954 | -34.5877 | -20.6263
12.5015 | -34.5881 | -20.6347
12.5076 | -34.5885 | -20.6434
12.5137 | -34.589 | -20.652
12.5198 | -34.5895 | -20.6604
12.5259 | -34.5899 | -20.669
12.532 | -34.5904 | -20.6775
12.5382 | -34.5908 | -20.6861
12.5443 | -34.5913 | -20.6946
12.5504 | -34.5917 | -20.7032
12.5565 | -34.5922 | -20.7116
12.5626 | -34.5927 | -20.7203
12.5687 | -34.5931 | -20.7289
12.5748 | -34.5935 | -20.7375
12.5809 | -34.594 | -20.746
12.587 | -34.5944 | -20.7547
12.5931 | -34.5949 | -20.7631
12.5992 | -34.5954 | -20.7716
12.6053 | -34.5958 | -20.7801
12.6114 | -34.5963 | -20.7884
12.6175 | -34.5967 | -20.797
12.6236 | -34.5972 | -20.8055
12.6297 | -34.5976 | -20.8139
12.6358 | -34.5981 | -20.8223
12.6419 | -34.5985 | -20.8308
12.648 | -34.599 | -20.8394
12.6541 | -34.5995 | -20.8479
12.6602 | -34.5999 | -20.8562
12.6663 | -34.6004 | -20.8647
12.6724 | -34.6008 | -20.8731
12.6785 | -34.6013 | -20.8815
12.6847 | -34.6017 | -20.8899
12.6908 | -34.6022 | -20.8985
12.6969 | -34.6027 | -20.9069
12.703 | -34.6031 | -20.9155
12.7091 | -34.6036 | -20.924
12.7152 | -34.604 | -20.9323
12.7213 | -34.6045 | -20.9408
12.7274 | -34.6049 | -20.9493
12.7335 | -34.6054 | -20.9575
12.7396 | -34.6059 | -20.9658
12.7457 | -34.6063 | -20.9743
12.7518 | -34.6068 | -20.9827
12.7579 | -34.6072 | -20.9913
12.764 | -34.6077 | -20.9998
12.7701 | -34.6081 | -21.0082
12.7762 | -34.6086 | -21.0167
12.7823 | -34.6091 | -21.0251
12.7884 | -34.6095 | -21.0337
12.7945 | -34.61 | -21.0423
12.8006 | -34.6105 | -21.0507
12.8067 | -34.6109 | -21.0592
12.8128 | -34.6114 | -21.0675
12.8189 | -34.6119 | -21.0759
12.8251 | -34.6123 | -21.0844
12.8312 | -34.6128 | -21.093
12.8373 | -34.6132 | -21.1014
12.8434 | -34.6137 | -21.1099
12.8495 | -34.6141 | -21.1183
12.8556 | -34.6146 | -21.1267
12.8617 | -34.615 | -21.1351
12.8678 | -34.6155 | -21.1437
12.8739 | -34.616 | -21.1521
12.88 | -34.6164 | -21.1606
12.8861 | -34.6168 | -21.169
12.8922 | -34.6173 | -21.1775
12.8983 | -34.6178 | -21.186
12.9044 | -34.6182 | -21.1943
12.9105 | -34.6187 | -21.2027
12.9166 | -34.6191 | -21.2111
12.9227 | -34.6196 | -21.2195
12.9288 | -34.62 | -21.2281
12.9349 | -34.6205 | -21.2366
12.941 | -34.621 | -21.245
12.9471 | -34.6214 | -21.2536
12.9532 | -34.6219 | -21.262
12.9593 | -34.6223 | -21.2706
12.9654 | -34.6228 | -21.2791
12.9716 | -34.6232 | -21.2878
12.9777 | -34.6237 | -21.2962
12.9838 | -34.6242 | -21.3045
12.9899 | -34.6246 | -21.313
12.996 | -34.625 | -21.3215
13.0021 | -34.6255 | -21.33
13.0082 | -34.626 | -21.3383
13.0143 | -34.6265 | -21.3466
13.0204 | -34.6269 | -21.3549
13.0265 | -34.6274 | -21.3633
13.0326 | -34.6278 | -21.3718
13.0387 | -34.6283 | -21.3804
13.0448 | -34.6287 | -21.3888
13.0509 | -34.6292 | -21.3972
13.057 | -34.6297 | -21.4055
13.0631 | -34.6302 | -21.4139
13.0692 | -34.6306 | -21.4222
13.0753 | -34.6311 | -21.4308
13.0814 | -34.6315 | -21.4394
13.0875 | -34.632 | -21.4477
13.0936 | -34.6324 | -21.4561
13.0997 | -34.6329 | -21.4645
13.1058 | -34.6334 | -21.4728
13.112 | -34.6338 | -21.4813
13.1181 | -34.6343 | -21.4896
13.1242 | -34.6348 | -21.4981
13.1303 | -34.6352 | -21.5064
13.1364 | -34.6357 | -21.5148
13.1425 | -34.6361 | -21.5232
13.1486 | -34.6366 | -21.5316
13.1547 | -34.637 | -21.54
13.1608 | -34.6375 | -21.5484
13.1669 | -34.6379 | -21.5567
13.173 | -34.6384 | -21.5652
13.1791 | -34.6389 | -21.5735
13.1852 | -34.6393 | -21.582
13.1913 | -34.6398 | -21.5904
13.1974 | -34.6402 | -21.5988
13.2035 | -34.6406 | -21.6072
13.2096 | -34.6411 | -21.6156
13.2157 | -34.6416 | -21.6238
13.2218 | -34.642 | -21.6321
13.2279 | -34.6425 | -21.6405
13.234 | -34.643 | -21.6488
13.2401 | -34.6434 | -21.6572
13.2462 | -34.6439 | -21.6655
13.2524 | -34.6443 | -21.6741
13.2585 | -34.6448 | -21.6824
13.2646 | -34.6453 | -21.6908
13.2707 | -34.6458 | -21.6992
13.2768 | -34.6462 | -21.7077
13.2829 | -34.6467 | -21.7162
13.289 | -34.6471 | -21.7244
13.2951 | -34.6476 | -21.7328
13.3012 | -34.6481 | -21.7412
13.3073 | -34.6485 | -21.7497
13.3134 | -34.649 | -21.7582
13.3195 | -34.6494 | -21.7665
13.3256 | -34.6499 | -21.7748
13.3317 | -34.6504 | -21.783
13.3378 | -34.6508 | -21.7914
13.3439 | -34.6513 | -21.7999
13.35 | -34.6518 | -21.8084
13.3561 | -34.6522 | -21.8166
13.3622 | -34.6527 | -21.8249
13.3683 | -34.6531 | -21.8332
13.3744 | -34.6536 | -21.8415
13.3805 | -34.6541 | -21.8499
13.3866 | -34.6545 | -21.8582
13.3927 | -34.6549 | -21.8666
13.3989 | -34.6554 | -21.8749
13.405 | -34.6559 | -21.8832
13.4111 | -34.6563 | -21.8915
13.4172 | -34.6568 | -21.8999
13.4233 | -34.6572 | -21.9084
13.4294 | -34.6577 | -21.9169
13.4355 | -34.6582 | -21.9251
13.4416 | -34.6586 | -21.9334
13.4477 | -34.6591 | -21.9418
13.4538 | -34.6596 | -21.9502
13.4599 | -34.66 | -21.9586
13.466 | -34.6605 | -21.967
13.4721 | -34.6609 | -21.9753
13.4782 | -34.6614 | -21.9837
13.4843 | -34.6618 | -21.9922
13.4904 | -34.6623 | -22.0005
13.4965 | -34.6627 | -22.009
13.5026 | -34.6632 | -22.0174
13.5087 | -34.6636 | -22.0257
13.5148 | -34.6641 | -22.0342
13.5209 | -34.6645 | -22.0427
13.527 | -34.665 | -22.051
13.5331 | -34.6655 | -22.0593
13.5393 | -34.6659 | -22.0675
13.5454 | -34.6664 | -22.0758
13.5515 | -34.6669 | -22.0843
13.5576 | -34.6673 | -22.0927
13.5637 | -34.6678 | -22.1011
13.5698 | -34.6682 | -22.1094
13.5759 | -34.6687 | -22.1175
13.582 | -34.6691 | -22.1259
13.5881 | -34.6696 | -22.1342
13.5942 | -34.6701 | -22.1427
13.6003 | -34.6705 | -22.151
13.6064 | -34.671 | -22.1596
13.6125 | -34.6715 | -22.168
13.6186 | -34.6719 | -22.1762
13.6247 | -34.6724 | -22.1846
13.6308 | -34.6728 | -22.1931
13.6369 | -34.6733 | -22.2015
13.643 | -34.6737 | -22.21
13.6491 | -34.6742 | -22.2182
13.6552 | -34.6747 | -22.2266
13.6613 | -34.6751 | -22.2349
13.6674 | -34.6756 | -22.2434
13.6735 | -34.676 | -22.2517
13.6796 | -34.6765 | -22.2601
13.6858 | -34.6769 | -22.2683
13.6919 | -34.6774 | -22.2766
13.698 | -34.6779 | -22.2851
13.7041 | -34.6783 | -22.2936
13.7102 | -34.6788 | -22.3019
13.7163 | -34.6793 | -22.3102
13.7224 | -34.6797 | -22.3185
13.7285 | -34.6802 | -22.3269
13.7346 | -34.6807 | -22.3353
13.7407 | -34.6811 | -22.3436
13.7468 | -34.6816 | -22.352
13.7529 | -34.682 | -22.3602
13.759 | -34.6825 | -22.3687
13.7651 | -34.6829 | -22.3771
13.7712 | -34.6834 | -22.3855
13.7773 | -34.6839 | -22.3939
13.7834 | -34.6843 | -22.4022
13.7895 | -34.6848 | -22.4106
13.7956 | -34.6852 | -22.419
13.8017 | -34.6857 | -22.4273
13.8078 | -34.6861 | -22.4356
13.8139 | -34.6866 | -22.4438
13.82 | -34.687 | -22.4521
13.8262 | -34.6875 | -22.4605
13.8323 | -34.6879 | -22.4689
13.8384 | -34.6884 | -22.4773
13.8445 | -34.6889 | -22.4856
13.8506 | -34.6893 | -22.494
13.8567 | -34.6898 | -22.5023
13.8628 | -34.6902 | -22.5107
13.8689 | -34.6907 | -22.5192
13.875 | -34.6912 | -22.5275
13.8811 | -34.6916 | -22.5358
13.8872 | -34.692 | -22.5442
13.8933 | -34.6925 | -22.5524
13.8994 | -34.6929 | -22.5607
13.9055 | -34.6934 | -22.5692
13.9116 | -34.6939 | -22.5775
13.9177 | -34.6943 | -22.5859
13.9238 | -34.6948 | -22.5943
13.9299 | -34.6952 | -22.6025
13.936 | -34.6957 | -22.6109
13.9421 | -34.6961 | -22.6193
13.9482 | -34.6966 | -22.6275
13.9543 | -34.697 | -22.636
13.9604 | -34.6975 | -22.6443
13.9665 | -34.6979 | -22.6526
13.9727 | -34.6984 | -22.6611
13.9788 | -34.6989 | -22.6693
13.9849 | -34.6993 | -22.6777
13.991 | -34.6998 | -22.686
13.9971 | -34.7002 | -22.6941
14.0032 | -34.7007 | -22.7024
14.0093 | -34.7011 | -22.7107
14.0154 | -34.7016 | -22.719
14.0215 | -34.7021 | -22.7272
14.0276 | -34.7025 | -22.7355
14.0337 | -34.703 | -22.7438
14.0398 | -34.7035 | -22.752
14.0459 | -34.7039 | -22.7603
14.052 | -34.7044 | -22.7687
14.0581 | -34.7048 | -22.777
14.0642 | -34.7052 | -22.7853
14.0703 | -34.7057 | -22.7936
14.0764 | -34.7062 | -22.8018
14.0825 | -34.7066 | -22.8101
14.0886 | -34.7071 | -22.8184
14.0947 | -34.7075 | -22.8268
14.1008 | -34.708 | -22.8351
14.1069 | -34.7085 | -22.8432
14.1131 | -34.7089 | -22.8516
14.1192 | -34.7094 | -22.8598
14.1253 | -34.7098 | -22.8681
14.1314 | -34.7103 | -22.8764
14.1375 | -34.7107 | -22.8847
14.1436 | -34.7112 | -22.8929
14.1497 | -34.7117 | -22.9012
14.1558 | -34.7121 | -22.9094
14.1619 | -34.7126 | -22.9176
14.168 | -34.713 | -22.926
14.1741 | -34.7135 | -22.9341
14.1802 | -34.7139 | -22.9424
14.1863 | -34.7144 | -22.9507
14.1924 | -34.7149 | -22.959
14.1985 | -34.7153 | -22.9672
14.2046 | -34.7158 | -22.9756
14.2107 | -34.7163 | -22.9838
14.2168 | -34.7167 | -22.9921
14.2229 | -34.7171 | -23.0004
14.229 | -34.7176 | -23.0088
14.2351 | -34.7181 | -23.017
14.2412 | -34.7185 | -23.0253
14.2473 | -34.719 | -23.0335
14.2534 | -34.7194 | -23.0416
14.2596 | -34.7199 | -23.0499
14.2657 | -34.7203 | -23.0581
14.2718 | -34.7208 | -23.0663
14.2779 | -34.7212 | -23.0746
14.284 | -34.7217 | -23.083
14.2901 | -34.7222 | -23.0912
14.2962 | -34.7226 | -23.0995
14.3023 | -34.7231 | -23.1079
14.3084 | -34.7235 | -23.1161
14.3145 | -34.724 | -23.1244
14.3206 | -34.7245 | -23.1328
14.3267 | -34.7249 | -23.1411
14.3328 | -34.7254 | -23.1493
14.3389 | -34.7258 | -23.1575
14.345 | -34.7263 | -23.1658
14.3511 | -34.7267 | -23.174
14.3572 | -34.7272 | -23.1824
14.3633 | -34.7277 | -23.1907
14.3694 | -34.7281 | -23.1989
14.3755 | -34.7286 | -23.2072
14.3816 | -34.7291 | -23.2156
14.3877 | -34.7295 | -23.2237
14.3938 | -34.73 | -23.2321
14.4 | -34.7304 | -23.2403
14.4061 | -34.7309 | -23.2485
14.4122 | -34.7313 | -23.2568
14.4183 | -34.7318 | -23.265
14.4244 | -34.7323 | -23.2732
14.4305 | -34.7327 | -23.2815
14.4366 | -34.7332 | -23.2897
14.4427 | -34.7337 | -23.2981
14.4488 | -34.7341 | -23.3064
14.4549 | -34.7346 | -23.3146
14.461 | -34.7351 | -23.3227
14.4671 | -34.7355 | -23.3311
14.4732 | -34.736 | -23.3393
14.4793 | -34.7364 | -23.3476
14.4854 | -34.7369 | -23.3558
14.4915 | -34.7373 | -23.3639
14.4976 | -34.7378 | -23.372
14.5037 | -34.7383 | -23.3801
14.5098 | -34.7387 | -23.3883
14.5159 | -34.7392 | -23.3966
14.522 | -34.7396 | -23.4048
14.5281 | -34.7401 | -23.4129
14.5342 | -34.7406 | -23.4212
14.5403 | -34.741 | -23.4293
14.5465 | -34.7415 | -23.4376
14.5526 | -34.7419 | -23.446
14.5587 | -34.7424 | -23.4542
14.5648 | -34.7429 | -23.4623
14.5709 | -34.7433 | -23.4705
14.577 | -34.7438 | -23.4786
14.5831 | -34.7442 | -23.4869
14.5892 | -34.7447 | -23.4951
14.5953 | -34.7451 | -23.5032
14.6014 | -34.7456 | -23.5114
14.6075 | -34.746 | -23.5195
14.6136 | -34.7465 | -23.5277
14.6197 | -34.7469 | -23.536
14.6258 | -34.7474 | -23.5442
14.6319 | -34.7478 | -23.5523
14.638 | -34.7483 | -23.5605
14.6441 | -34.7487 | -23.5687
14.6502 | -34.7492 | -23.577
14.6563 | -34.7497 | -23.5852
14.6624 | -34.7501 | -23.5934
14.6685 | -34.7506 | -23.6014
14.6746 | -34.7511 | -23.6096
14.6807 | -34.7515 | -23.6177
14.6869 | -34.752 | -23.6259
14.693 | -34.7525 | -23.634
14.6991 | -34.7529 | -23.6422
14.7052 | -34.7534 | -23.6502
14.7113 | -34.7539 | -23.6585
14.7174 | -34.7543 | -23.6667
14.7235 | -34.7548 | -23.6748
14.7296 | -34.7552 | -23.6829
14.7357 | -34.7557 | -23.691
14.7418 | -34.7562 | -23.6992
14.7479 | -34.7566 | -23.7073
14.754 | -34.7571 | -23.7155
14.7601 | -34.7576 | -23.7236
14.7662 | -34.758 | -23.7317
14.7723 | -34.7585 | -23.7396
14.7784 | -34.7589 | -23.7478
14.7845 | -34.7594 | -23.756
14.7906 | -34.7599 | -23.7642
14.7967 | -34.7603 | -23.7724
14.8028 | -34.7608 | -23.7805
14.8089 | -34.7612 | -23.7886
14.815 | -34.7617 | -23.7967
14.8211 | -34.7622 | -23.8049
14.8272 | -34.7626 | -23.8132
14.8334 | -34.7631 | -23.8213
14.8395 | -34.7636 | -23.8294
14.8456 | -34.764 | -23.8374
14.8517 | -34.7645 | -23.8456
14.8578 | -34.7649 | -23.8537
14.8639 | -34.7654 | -23.862
14.87 | -34.7659 | -23.8702
14.8761 | -34.7664 | -23.8782
14.8822 | -34.7668 | -23.8864
14.8883 | -34.7673 | -23.8946
14.8944 | -34.7677 | -23.9027
14.9005 | -34.7682 | -23.9108
14.9066 | -34.7687 | -23.919
14.9127 | -34.7691 | -23.927
14.9188 | -34.7696 | -23.9351
14.9249 | -34.77 | -23.9432
14.931 | -34.7705 | -23.9514
14.9371 | -34.771 | -23.9594
14.9432 | -34.7714 | -23.9676
14.9493 | -34.7719 | -23.9756
14.9554 | -34.7724 | -23.9837
14.9615 | -34.7728 | -23.9919
14.9676 | -34.7733 | -24.0001
14.9738 | -34.7738 | -24.0081
14.9799 | -34.7742 | -24.0164
14.986 | -34.7747 | -24.0244
14.9921 | -34.7751 | -24.0325
14.9982 | -34.7756 | -24.0406
15.0043 | -34.776 | -24.0486
15.0104 | -34.7765 | -24.0568
15.0165 | -34.777 | -24.065
15.0226 | -34.7775 | -24.0729
15.0287 | -34.7779 | -24.0811
15.0348 | -34.7784 | -24.0892
15.0409 | -34.7789 | -24.0974
15.047 | -34.7793 | -24.1056
15.0531 | -34.7798 | -24.1136
15.0592 | -34.7803 | -24.1217
15.0653 | -34.7808 | -24.1299
15.0714 | -34.7812 | -24.138
15.0775 | -34.7817 | -24.146
15.0836 | -34.7822 | -24.1541
15.0897 | -34.7827 | -24.1619
15.0958 | -34.7831 | -24.17
15.1019 | -34.7836 | -24.1781
15.108 | -34.7841 | -24.1862
15.1141 | -34.7845 | -24.1943
15.1203 | -34.785 | -24.2024
15.1264 | -34.7855 | -24.2104
15.1325 | -34.786 | -24.2186
15.1386 | -34.7864 | -24.2266
15.1447 | -34.7869 | -24.2347
15.1508 | -34.7874 | -24.2428
15.1569 | -34.7879 | -24.2509
15.163 | -34.7883 | -24.2587
15.1691 | -34.7888 | -24.2667
15.1752 | -34.7893 | -24.2747
15.1813 | -34.7897 | -24.2827
15.1874 | -34.7902 | -24.2908
15.1935 | -34.7907 | -24.2987
15.1996 | -34.7911 | -24.3067
15.2057 | -34.7916 | -24.3147
15.2118 | -34.7921 | -24.3227
15.2179 | -34.7925 | -24.3309
15.224 | -34.793 | -24.3389
15.2301 | -34.7935 | -24.3467
15.2362 | -34.794 | -24.3548
15.2423 | -34.7944 | -24.3629
15.2484 | -34.7949 | -24.371
15.2545 | -34.7954 | -24.3791
15.2607 | -34.7958 | -24.3871
15.2668 | -34.7963 | -24.3948
15.2729 | -34.7968 | -24.4028
15.279 | -34.7973 | -24.4109
15.2851 | -34.7978 | -24.4189
15.2912 | -34.7983 | -24.4269
15.2973 | -34.7987 | -24.4351
15.3034 | -34.7992 | -24.443
15.3095 | -34.7997 | -24.451
15.3156 | -34.8002 | -24.459
15.3217 | -34.8007 | -24.467
15.3278 | -34.8011 | -24.4752
15.3339 | -34.8017 | -24.483
15.34 | -34.8021 | -24.4909
15.3461 | -34.8026 | -24.4989
15.3522 | -34.8031 | -24.5068
15.3583 | -34.8035 | -24.515
15.3644 | -34.804 | -24.5229
15.3705 | -34.8045 | -24.5306
15.3766 | -34.805 | -24.5388
15.3827 | -34.8054 | -24.5466
15.3888 | -34.8059 | -24.5547
15.3949 | -34.8064 | -24.5627
15.401 | -34.8069 | -24.5708
15.4072 | -34.8074 | -24.5785
15.4133 | -34.8078 | -24.5865
15.4194 | -34.8083 | -24.5946
15.4255 | -34.8088 | -24.6026
15.4316 | -34.8092 | -24.6106
15.4377 | -34.8097 | -24.6188
15.4438 | -34.8102 | -24.6266
15.4499 | -34.8107 | -24.6346
15.456 | -34.8112 | -24.6424
15.4621 | -34.8117 | -24.6504
15.4682 | -34.8121 | -24.6585
15.4743 | -34.8126 | -24.6663
15.4804 | -34.8131 | -24.6742
15.4865 | -34.8136 | -24.6822
15.4926 | -34.814 | -24.6902
15.4987 | -34.8145 | -24.6981
15.5048 | -34.815 | -24.7061
15.5109 | -34.8155 | -24.7139
15.517 | -34.816 | -24.7219
15.5231 | -34.8165 | -24.7298
15.5292 | -34.8169 | -24.7377
15.5353 | -34.8174 | -24.7458
15.5414 | -34.8179 | -24.7537
15.5476 | -34.8184 | -24.7614
15.5537 | -34.8188 | -24.7694
15.5598 | -34.8193 | -24.7774
15.5659 | -34.8198 | -24.7854
15.572 | -34.8203 | -24.7933
15.5781 | -34.8208 | -24.8011
15.5842 | -34.8213 | -24.8089
15.5903 | -34.8217 | -24.8167
15.5964 | -34.8222 | -24.8245
15.6025 | -34.8227 | -24.8323
15.6086 | -34.8232 | -24.8403
15.6147 | -34.8237 | -24.8483
15.6208 | -34.8241 | -24.856
15.6269 | -34.8246 | -24.8638
15.633 | -34.8251 | -24.8718
15.6391 | -34.8256 | -24.8798
15.6452 | -34.8261 | -24.8878
15.6513 | -34.8266 | -24.8954
15.6574 | -34.8271 | -24.9034
15.6635 | -34.8276 | -24.9114
15.6696 | -34.8281 | -24.9193
15.6757 | -34.8285 | -24.9273
15.6818 | -34.829 | -24.9352
15.688 | -34.8295 | -24.943
15.6941 | -34.83 | -24.9508
15.7002 | -34.8304 | -24.9587
15.7063 | -34.8309 | -24.9666
15.7124 | -34.8314 | -24.9746
15.7185 | -34.8318 | -24.9825
15.7246 | -34.8323 | -24.9902
15.7307 | -34.8328 | -24.998
15.7368 | -34.8333 | -25.0059
15.7429 | -34.8338 | -25.0138
15.749 | -34.8343 | -25.0218
15.7551 | -34.8348 | -25.0297
15.7612 | -34.8353 | -25.0374
15.7673 | -34.8358 | -25.0452
15.7734 | -34.8363 | -25.0531
15.7795 | -34.8368 | -25.061
15.7856 | -34.8372 | -25.0688
15.7917 | -34.8378 | -25.0764
15.7978 | -34.8382 | -25.0843
15.8039 | -34.8387 | -25.0922
15.81 | -34.8392 | -25.1001
15.8161 | -34.8397 | -25.1079
15.8222 | -34.8402 | -25.1157
15.8283 | -34.8407 | -25.1233
15.8345 | -34.8411 | -25.1311
15.8406 | -34.8416 | -25.1389
15.8467 | -34.8421 | -25.1469
15.8528 | -34.8426 | -25.1547
15.8589 | -34.8431 | -25.1625
15.865 | -34.8436 | -25.17
15.8711 | -34.8441 | -25.1779
15.8772 | -34.8446 | -25.1857
15.8833 | -34.8451 | -25.1936
15.8894 | -34.8456 | -25.2015
15.8955 | -34.846 | -25.2092
15.9016 | -34.8465 | -25.2168
15.9077 | -34.847 | -25.2246
15.9138 | -34.8475 | -25.2325
15.9199 | -34.848 | -25.2403
15.926 | -34.8484 | -25.248
15.9321 | -34.8489 | -25.2557
15.9382 | -34.8494 | -25.2635
15.9443 | -34.8499 | -25.2713
15.9504 | -34.8504 | -25.2792
15.9565 | -34.8509 | -25.287
15.9626 | -34.8513 | -25.2949
15.9687 | -34.8518 | -25.3026
15.9749 | -34.8523 | -25.3103
15.981 | -34.8528 | -25.3181
15.9871 | -34.8533 | -25.326
15.9932 | -34.8538 | -25.3337
15.9993 | -34.8542 | -25.3415
16.0054 | -34.8548 | -25.3491
16.0115 | -34.8552 | -25.357
16.0176 | -34.8557 | -25.3648
16.0237 | -34.8562 | -25.3727
16.0298 | -34.8567 | -25.3804
16.0359 | -34.8572 | -25.3881
16.042 | -34.8577 | -25.3957
16.0481 | -34.8582 | -25.4036
16.0542 | -34.8587 | -25.4114
16.0603 | -34.8592 | -25.4191
16.0664 | -34.8596 | -25.4268
16.0725 | -34.8601 | -25.4343
16.0786 | -34.8606 | -25.442
16.0847 | -34.8611 | -25.4499
16.0908 | -34.8616 | -25.4577
16.0969 | -34.8621 | -25.4655
16.103 | -34.8625 | -25.4733
16.1091 | -34.8631 | -25.4808
16.1152 | -34.8635 | -25.4885
16.1214 | -34.864 | -25.4963
16.1275 | -34.8645 | -25.5041
16.1336 | -34.865 | -25.5117
16.1397 | -34.8655 | -25.5194
16.1458 | -34.866 | -25.5269
16.1519 | -34.8665 | -25.5348
16.158 | -34.867 | -25.5425
16.1641 | -34.8674 | -25.5504
16.1702 | -34.8679 | -25.5582
16.1763 | -34.8684 | -25.566
16.1824 | -34.8689 | -25.5735
16.1885 | -34.8694 | -25.5811
16.1946 | -34.8699 | -25.5888
16.2007 | -34.8704 | -25.5966
16.2068 | -34.8709 | -25.6043
16.2129 | -34.8714 | -25.612
16.219 | -34.8719 | -25.6195
16.2251 | -34.8724 | -25.6271
16.2312 | -34.8729 | -25.6348
16.2373 | -34.8733 | -25.6424
16.2434 | -34.8738 | -25.6501
16.2495 | -34.8743 | -25.6575
16.2556 | -34.8748 | -25.6653
16.2618 | -34.8753 | -25.6729
16.2679 | -34.8758 | -25.6806
16.274 | -34.8763 | -25.6883
16.2801 | -34.8768 | -25.6959
16.2862 | -34.8773 | -25.7034
16.2923 | -34.8778 | -25.7112
16.2984 | -34.8783 | -25.7188
16.3045 | -34.8787 | -25.7265
16.3106 | -34.8792 | -25.7342
16.3167 | -34.8797 | -25.7419
16.3228 | -34.8802 | -25.7495
16.3289 | -34.8807 | -25.7572
16.335 | -34.8812 | -25.7647
16.3411 | -34.8817 | -25.7724
16.3472 | -34.8822 | -25.78
16.3533 | -34.8827 | -25.7877
16.3594 | -34.8831 | -25.7951
16.3655 | -34.8836 | -25.8028
16.3716 | -34.8841 | -25.8104
16.3777 | -34.8846 | -25.818
16.3838 | -34.8851 | -25.8257
16.3899 | -34.8856 | -25.8333
16.396 | -34.8861 | -25.841
16.4021 | -34.8865 | -25.8486
16.4083 | -34.887 | -25.8562
16.4144 | -34.8875 | -25.8638
16.4205 | -34.888 | -25.8715
16.4266 | -34.8885 | -25.879
16.4327 | -34.889 | -25.8865
16.4388 | -34.8895 | -25.8942
16.4449 | -34.89 | -25.9019
16.451 | -34.8904 | -25.9096
16.4571 | -34.8909 | -25.9174
16.4632 | -34.8914 | -25.9247
16.4693 | -34.8919 | -25.9323
16.4754 | -34.8924 | -25.9401
16.4815 | -34.8929 | -25.9478
16.4876 | -34.8933 | -25.9553
16.4937 | -34.8938 | -25.9629
16.4998 | -34.8943 | -25.9703
16.5059 | -34.8948 | -25.978
16.512 | -34.8953 | -25.9857
16.5181 | -34.8958 | -25.9934
16.5242 | -34.8963 | -26.0011
16.5303 | -34.8968 | -26.0087
16.5364 | -34.8972 | -26.0163
16.5425 | -34.8977 | -26.0239
16.5487 | -34.8982 | -26.0314
16.5548 | -34.8987 | -26.0391
16.5609 | -34.8992 | -26.0468
16.567 | -34.8997 | -26.0542
16.5731 | -34.9002 | -26.0619
16.5792 | -34.9007 | -26.0694
16.5853 | -34.9012 | -26.077
16.5914 | -34.9017 | -26.0847
16.5975 | -34.9021 | -26.0923
16.6036 | -34.9026 | -26.0997
16.6097 | -34.9031 | -26.1073
16.6158 | -34.9036 | -26.1148
16.6219 | -34.9041 | -26.1225
16.628 | -34.9046 | -26.13
16.6341 | -34.9051 | -26.1377
16.6402 | -34.9055 | -26.1452
16.6463 | -34.906 | -26.1527
16.6524 | -34.9065 | -26.1602
16.6585 | -34.907 | -26.1679
16.6646 | -34.9075 | -26.1754
16.6707 | -34.908 | -26.1829
16.6768 | -34.9085 | -26.1905
16.6829 | -34.909 | -26.1981
16.689 | -34.9094 | -26.2057
16.6952 | -34.9099 | -26.2133
16.7013 | -34.9104 | -26.2207
16.7074 | -34.9109 | -26.2281
16.7135 | -34.9114 | -26.2357
16.7196 | -34.9118 | -26.2433
16.7257 | -34.9123 | -26.2508
16.7318 | -34.9128 | -26.2584
16.7379 | -34.9133 | -26.2658
16.744 | -34.9138 | -26.2732
16.7501 | -34.9143 | -26.2809
16.7562 | -34.9148 | -26.2885
16.7623 | -34.9152 | -26.2959
16.7684 | -34.9157 | -26.3035
16.7745 | -34.9162 | -26.3112
16.7806 | -34.9167 | -26.3185
16.7867 | -34.9172 | -26.326
16.7928 | -34.9177 | -26.3336
16.7989 | -34.9182 | -26.3411
16.805 | -34.9186 | -26.3487
16.8111 | -34.9191 | -26.356
16.8172 | -34.9196 | -26.3636
16.8233 | -34.9201 | -26.3712
16.8294 | -34.9206 | -26.3787
16.8356 | -34.9211 | -26.3863
16.8417 | -34.9216 | -26.3939
16.8478 | -34.9221 | -26.4013
16.8539 | -34.9226 | -26.4087
16.86 | -34.923 | -26.4164
16.8661 | -34.9235 | -26.4238
16.8722 | -34.924 | -26.4314
16.8783 | -34.9245 | -26.439
16.8844 | -34.925 | -26.4464
16.8905 | -34.9255 | -26.4539
16.8966 | -34.9259 | -26.4613
16.9027 | -34.9264 | -26.4689
16.9088 | -34.9269 | -26.4766
16.9149 | -34.9274 | -26.4841
16.921 | -34.9279 | -26.4914
16.9271 | -34.9284 | -26.499
16.9332 | -34.9289 | -26.5066
16.9393 | -34.9294 | -26.5141
16.9454 | -34.9299 | -26.5216
16.9515 | -34.9303 | -26.5292
16.9576 | -34.9308 | -26.5365
16.9637 | -34.9313 | -26.544
16.9698 | -34.9318 | -26.5513
16.9759 | -34.9323 | -26.5589
16.9821 | -34.9328 | -26.5664
16.9882 | -34.9332 | -26.5738
16.9943 | -34.9337 | -26.5812
17.0004 | -34.9342 | -26.5887
17.0065 | -34.9347 | -26.5962
17.0126 | -34.9352 | -26.6037
17.0187 | -34.9357 | -26.6114
17.0248 | -34.9362 | -26.6188
17.0309 | -34.9367 | -26.6263
17.037 | -34.9371 | -26.6337
17.0431 | -34.9376 | -26.6412
17.0492 | -34.9381 | -26.6487
17.0553 | -34.9386 | -26.6562
17.0614 | -34.9391 | -26.6635
17.0675 | -34.9395 | -26.671
17.0736 | -34.94 | -26.6785
17.0797 | -34.9405 | -26.686
17.0858 | -34.941 | -26.6935
17.0919 | -34.9415 | -26.701
17.098 | -34.942 | -26.7084
17.1041 | -34.9424 | -26.7158
17.1102 | -34.9429 | -26.7232
17.1163 | -34.9434 | -26.7306
17.1225 | -34.9439 | -26.7379
17.1286 | -34.9444 | -26.7452
17.1347 | -34.9449 | -26.7526
17.1408 | -34.9454 | -26.7602
17.1469 | -34.9459 | -26.7677
17.153 | -34.9464 | -26.7752
17.1591 | -34.9468 | -26.7828
17.1652 | -34.9473 | -26.79
17.1713 | -34.9478 | -26.7974
17.1774 | -34.9483 | -26.8048
17.1835 | -34.9488 | -26.8121
17.1896 | -34.9492 | -26.8197
17.1957 | -34.9497 | -26.8271
17.2018 | -34.9502 | -26.8342
17.2079 | -34.9507 | -26.8417
17.214 | -34.9512 | -26.8493
17.2201 | -34.9517 | -26.8568
17.2262 | -34.9521 | -26.8643
17.2323 | -34.9526 | -26.8718
17.2384 | -34.9531 | -26.8792
17.2445 | -34.9536 | -26.8865
17.2506 | -34.9541 | -26.894
17.2567 | -34.9546 | -26.9015
17.2628 | -34.9551 | -26.9089
17.269 | -34.9556 | -26.9162
17.2751 | -34.9561 | -26.9237
17.2812 | -34.9565 | -26.9311
17.2873 | -34.957 | -26.9385
17.2934 | -34.9575 | -26.946
17.2995 | -34.958 | -26.9535
17.3056 | -34.9585 | -26.9608
17.3117 | -34.959 | -26.9683
17.3178 | -34.9595 | -26.9756
17.3239 | -34.9599 | -26.983
17.33 | -34.9604 | -26.9904
17.3361 | -34.9609 | -26.9978
17.3422 | -34.9613 | -27.0049
17.3483 | -34.9618 | -27.0124
17.3544 | -34.9623 | -27.0199
17.3605 | -34.9628 | -27.0273
17.3666 | -34.9633 | -27.0348
17.3727 | -34.9637 | -27.0422
17.3788 | -34.9642 | -27.0495
17.3849 | -34.9647 | -27.0569
17.391 | -34.9652 | -27.0642
17.3971 | -34.9657 | -27.0716
17.4032 | -34.9661 | -27.0791
17.4094 | -34.9666 | -27.0863
17.4155 | -34.9671 | -27.0936
17.4216 | -34.9676 | -27.1009
17.4277 | -34.9681 | -27.1083
17.4338 | -34.9686 | -27.1158
17.4399 | -34.9691 | -27.1232
17.446 | -34.9695 | -27.1304
17.4521 | -34.97 | -27.1378
17.4582 | -34.9705 | -27.1451
17.4643 | -34.971 | -27.1524
17.4704 | -34.9715 | -27.1598
17.4765 | -34.9719 | -27.1673
17.4826 | -34.9724 | -27.1746
17.4887 | -34.9729 | -27.1821
17.4948 | -34.9734 | -27.1894
17.5009 | -34.9739 | -27.1968
17.507 | -34.9743 | -27.2043
17.5131 | -34.9748 | -27.2116
17.5192 | -34.9753 | -27.2189
17.5253 | -34.9758 | -27.2263
17.5314 | -34.9763 | -27.2337
17.5375 | -34.9768 | -27.2409
17.5436 | -34.9772 | -27.2483
17.5497 | -34.9777 | -27.2557
17.5559 | -34.9782 | -27.2629
17.562 | -34.9787 | -27.2702
17.5681 | -34.9792 | -27.2775
17.5742 | -34.9797 | -27.285
17.5803 | -34.9802 | -27.2923
17.5864 | -34.9807 | -27.2995
17.5925 | -34.9811 | -27.3069
17.5986 | -34.9816 | -27.3142
17.6047 | -34.9821 | -27.3217
17.6108 | -34.9826 | -27.329
17.6169 | -34.9831 | -27.3364
17.623 | -34.9836 | -27.3436
17.6291 | -34.9841 | -27.3509
17.6352 | -34.9845 | -27.3581
17.6413 | -34.985 | -27.3656
17.6474 | -34.9855 | -27.3731
17.6535 | -34.986 | -27.3804
17.6596 | -34.9865 | -27.3876
17.6657 | -34.987 | -27.395
17.6718 | -34.9875 | -27.4024
17.6779 | -34.9879 | -27.4096
17.684 | -34.9884 | -27.4169
17.6901 | -34.9889 | -27.4243
17.6963 | -34.9894 | -27.4315
17.7024 | -34.9899 | -27.4388
17.7085 | -34.9903 | -27.446
17.7146 | -34.9908 | -27.4532
17.7207 | -34.9913 | -27.4604
17.7268 | -34.9918 | -27.4677
17.7329 | -34.9923 | -27.475
17.739 | -34.9928 | -27.4824
17.7451 | -34.9932 | -27.4896
17.7512 | -34.9937 | -27.497
17.7573 | -34.9942 | -27.5045
17.7634 | -34.9946 | -27.5117
17.7695 | -34.9951 | -27.519
17.7756 | -34.9956 | -27.5263
17.7817 | -34.9961 | -27.5336
17.7878 | -34.9965 | -27.541
17.7939 | -34.997 | -27.5484
17.8 | -34.9975 | -27.5556
17.8061 | -34.998 | -27.5629
17.8122 | -34.9985 | -27.5703
17.8183 | -34.9989 | -27.5776
17.8244 | -34.9994 | -27.5849
17.8305 | -34.9999 | -27.5923
17.8366 | -35.0004 | -27.5995
17.8428 | -35.0009 | -27.6068
17.8489 | -35.0013 | -27.614
17.855 | -35.0018 | -27.6213
17.8611 | -35.0023 | -27.6286
17.8672 | -35.0028 | -27.6356
17.8733 | -35.0033 | -27.6431
17.8794 | -35.0038 | -27.6503
17.8855 | -35.0042 | -27.6576
17.8916 | -35.0047 | -27.665
17.8977 | -35.0052 | -27.6723
17.9038 | -35.0057 | -27.6794
17.9099 | -35.0061 | -27.6867
17.916 | -35.0066 | -27.6941
17.9221 | -35.0071 | -27.7015
17.9282 | -35.0076 | -27.7088
17.9343 | -35.0081 | -27.7162
17.9404 | -35.0086 | -27.7233
17.9465 | -35.0091 | -27.7307
17.9526 | -35.0095 | -27.738
17.9587 | -35.01 | -27.7452
17.9648 | -35.0105 | -27.7524
17.9709 | -35.0109 | -27.7597
17.977 | -35.0114 | -27.767
17.9832 | -35.0119 | -27.7744
17.9893 | -35.0123 | -27.7817
17.9954 | -35.0128 | -27.789
18.0015 | -35.0133 | -27.7964
18.0076 | -35.0138 | -27.8036
18.0137 | -35.0143 | -27.8109
18.0198 | -35.0147 | -27.8181
18.0259 | -35.0152 | -27.8254
18.032 | -35.0157 | -27.8327
18.0381 | -35.0161 | -27.8401
18.0442 | -35.0166 | -27.8472
18.0503 | -35.0171 | -27.8545
18.0564 | -35.0176 | -27.8617
18.0625 | -35.0181 | -27.869
18.0686 | -35.0185 | -27.8764
18.0747 | -35.019 | -27.8838
18.0808 | -35.0195 | -27.891
18.0869 | -35.02 | -27.8984
18.093 | -35.0204 | -27.9057
18.0991 | -35.0209 | -27.9131
18.1052 | -35.0214 | -27.9203
18.1113 | -35.0218 | -27.9276
18.1174 | -35.0223 | -27.9347
18.1236 | -35.0228 | -27.9421
18.1297 | -35.0233 | -27.9494
18.1358 | -35.0238 | -27.9567
18.1419 | -35.0243 | -27.9639
18.148 | -35.0247 | -27.971
18.1541 | -35.0252 | -27.9779
18.1602 | -35.0257 | -27.9851
18.1663 | -35.0262 | -27.9922
18.1724 | -35.0267 | -27.9994
18.1785 | -35.0271 | -28.0065
18.1846 | -35.0276 | -28.0137
18.1907 | -35.028 | -28.0211
18.1968 | -35.0285 | -28.0284
18.2029 | -35.029 | -28.0356
18.209 | -35.0295 | -28.0429
18.2151 | -35.0299 | -28.0503
18.2212 | -35.0304 | -28.0574
18.2273 | -35.0309 | -28.0646
18.2334 | -35.0314 | -28.0718
18.2395 | -35.0318 | -28.0791
18.2456 | -35.0323 | -28.0863
18.2517 | -35.0328 | -28.0937
18.2578 | -35.0332 | -28.1009
18.2639 | -35.0337 | -28.1082
18.2701 | -35.0342 | -28.1154
18.2762 | -35.0347 | -28.1227
18.2823 | -35.0352 | -28.1298
18.2884 | -35.0357 | -28.1369
18.2945 | -35.0362 | -28.144
18.3006 | -35.0367 | -28.1514
18.3067 | -35.0372 | -28.1586
18.3128 | -35.0376 | -28.1659
18.3189 | -35.0381 | -28.1732
18.325 | -35.0386 | -28.1805
18.3311 | -35.0391 | -28.1876
18.3372 | -35.0396 | -28.195
18.3433 | -35.04 | -28.2023
18.3494 | -35.0405 | -28.2095
18.3555 | -35.041 | -28.2168
18.3616 | -35.0415 | -28.2239
18.3677 | -35.042 | -28.2312
18.3738 | -35.0425 | -28.2383
18.3799 | -35.043 | -28.2454
18.386 | -35.0434 | -28.2526
18.3921 | -35.0439 | -28.2598
18.3982 | -35.0444 | -28.2669
18.4043 | -35.0449 | -28.274
18.4105 | -35.0453 | -28.2813
18.4166 | -35.0458 | -28.2885
18.4227 | -35.0463 | -28.2957
18.4288 | -35.0468 | -28.3029
18.4349 | -35.0473 | -28.31
18.441 | -35.0477 | -28.3171
18.4471 | -35.0482 | -28.3243
18.4532 | -35.0487 | -28.3316
18.4593 | -35.0491 | -28.3388
18.4654 | -35.0496 | -28.3457
18.4715 | -35.0501 | -28.3531
18.4776 | -35.0505 | -28.3604
18.4837 | -35.051 | -28.3678
18.4898 | -35.0515 | -28.3751
18.4959 | -35.052 | -28.3824
18.502 | -35.0525 | -28.3894
18.5081 | -35.0529 | -28.3967
18.5142 | -35.0534 | -28.4039
18.5203 | -35.0539 | -28.4111
18.5264 | -35.0543 | -28.4181
18.5325 | -35.0548 | -28.4254
18.5386 | -35.0553 | -28.4326
18.5447 | -35.0558 | -28.4396
18.5508 | -35.0562 | -28.4467
18.557 | -35.0567 | -28.4541
18.5631 | -35.0572 | -28.4613
18.5692 | -35.0577 | -28.4684
18.5753 | -35.0582 | -28.4755
18.5814 | -35.0586 | -28.4825
18.5875 | -35.0591 | -28.4898
18.5936 | -35.0596 | -28.497
18.5997 | -35.06 | -28.5043
18.6058 | -35.0605 | -28.5113
18.6119 | -35.061 | -28.5186
18.618 | -35.0614 | -28.5257
18.6241 | -35.0619 | -28.5329
18.6302 | -35.0624 | -28.5402
18.6363 | -35.0629 | -28.5475
18.6424 | -35.0634 | -28.5547
18.6485 | -35.0638 | -28.562
18.6546 | -35.0643 | -28.5691
18.6607 | -35.0648 | -28.5764
18.6668 | -35.0653 | -28.5836
18.6729 | -35.0658 | -28.5907
18.679 | -35.0663 | -28.5977
18.6851 | -35.0667 | -28.6048
18.6912 | -35.0672 | -28.6119
18.6974 | -35.0676 | -28.619
18.7035 | -35.0681 | -28.6262
18.7096 | -35.0686 | -28.6335
18.7157 | -35.0691 | -28.6405
18.7218 | -35.0695 | -28.6479
18.7279 | -35.07 | -28.655
18.734 | -35.0705 | -28.6622
18.7401 | -35.071 | -28.6692
18.7462 | -35.0715 | -28.6763
18.7523 | -35.072 | -28.6835
18.7584 | -35.0725 | -28.6908
18.7645 | -35.073 | -28.6979
18.7706 | -35.0734 | -28.7051
18.7767 | -35.0739 | -28.7122
18.7828 | -35.0744 | -28.7192
18.7889 | -35.0748 | -28.7264
18.795 | -35.0753 | -28.7336
18.8011 | -35.0758 | -28.7408
18.8072 | -35.0763 | -28.748
18.8133 | -35.0767 | -28.7552
18.8194 | -35.0772 | -28.7621
18.8255 | -35.0777 | -28.7693
18.8316 | -35.0781 | -28.7764
18.8377 | -35.0786 | -28.7835
18.8439 | -35.0791 | -28.7907
18.85 | -35.0795 | -28.7979
18.8561 | -35.08 | -28.805
18.8622 | -35.0805 | -28.8122
18.8683 | -35.0809 | -28.8194
18.8744 | -35.0814 | -28.8266
18.8805 | -35.0819 | -28.8338
18.8866 | -35.0824 | -28.8411
18.8927 | -35.0829 | -28.8482
18.8988 | -35.0834 | -28.8553
18.9049 | -35.0838 | -28.8624
18.911 | -35.0843 | -28.8694
18.9171 | -35.0848 | -28.8764
18.9232 | -35.0853 | -28.8835
18.9293 | -35.0858 | -28.8907
18.9354 | -35.0862 | -28.8977
18.9415 | -35.0867 | -28.9048
18.9476 | -35.0872 | -28.9119
18.9537 | -35.0876 | -28.9192
18.9598 | -35.0881 | -28.926
18.9659 | -35.0886 | -28.9332
18.972 | -35.0891 | -28.9402
18.9781 | -35.0895 | -28.9473
18.9843 | -35.09 | -28.9545
18.9904 | -35.0905 | -28.9616
18.9965 | -35.091 | -28.9686
19.0026 | -35.0914 | -28.9758
19.0087 | -35.0919 | -28.983
19.0148 | -35.0924 | -28.9901
19.0209 | -35.0929 | -28.9973
19.027 | -35.0933 | -29.0045
19.0331 | -35.0938 | -29.0116
19.0392 | -35.0943 | -29.0186
19.0453 | -35.0948 | -29.0257
19.0514 | -35.0953 | -29.0328
19.0575 | -35.0957 | -29.0401
19.0636 | -35.0962 | -29.0468
19.0697 | -35.0967 | -29.0539
19.0758 | -35.0971 | -29.0609
19.0819 | -35.0976 | -29.068
19.088 | -35.0981 | -29.0751
19.0941 | -35.0985 | -29.0823
19.1002 | -35.099 | -29.0892
19.1063 | -35.0995 | -29.0963
19.1124 | -35.1 | -29.1034
19.1185 | -35.1004 | -29.1105
19.1246 | -35.1009 | -29.1174
19.1308 | -35.1014 | -29.1245
19.1369 | -35.1019 | -29.1314
19.143 | -35.1024 | -29.1386
19.1491 | -35.1029 | -29.1457
19.1552 | -35.1033 | -29.1528
19.1613 | -35.1038 | -29.16
19.1674 | -35.1043 | -29.1671
19.1735 | -35.1047 | -29.1741
19.1796 | -35.1052 | -29.1812
19.1857 | -35.1057 | -29.1883
19.1918 | -35.1062 | -29.1955
19.1979 | -35.1066 | -29.2024
19.204 | -35.1071 | -29.2091
19.2101 | -35.1076 | -29.2161
19.2162 | -35.108 | -29.2231
19.2223 | -35.1085 | -29.23
19.2284 | -35.109 | -29.2373
19.2345 | -35.1095 | -29.2445
19.2406 | -35.1099 | -29.2516
19.2467 | -35.1104 | -29.2588
19.2528 | -35.1109 | -29.2659
19.2589 | -35.1113 | -29.2728
19.265 | -35.1118 | -29.2801
19.2712 | -35.1123 | -29.2871
19.2773 | -35.1127 | -29.294
19.2834 | -35.1132 | -29.3009
19.2895 | -35.1137 | -29.308
19.2956 | -35.1142 | -29.3152
19.3017 | -35.1146 | -29.3222
19.3078 | -35.1151 | -29.3291
19.3139 | -35.1156 | -29.3359
19.32 | -35.1161 | -29.3428
19.3261 | -35.1165 | -29.3499
19.3322 | -35.117 | -29.3569
19.3383 | -35.1175 | -29.3637
19.3444 | -35.118 | -29.3705
19.3505 | -35.1185 | -29.3776
19.3566 | -35.1189 | -29.3845
19.3627 | -35.1194 | -29.3915
19.3688 | -35.1199 | -29.3984
19.3749 | -35.1204 | -29.4055
19.381 | -35.1208 | -29.4123
19.3871 | -35.1213 | -29.419
19.3932 | -35.1218 | -29.4262
19.3993 | -35.1222 | -29.4332
19.4054 | -35.1227 | -29.4402
19.4115 | -35.1231 | -29.4471
19.4177 | -35.1236 | -29.4538
19.4238 | -35.1241 | -29.4608
19.4299 | -35.1245 | -29.4678
19.436 | -35.125 | -29.4749
19.4421 | -35.1254 | -29.4819
19.4482 | -35.1259 | -29.4889
19.4543 | -35.1264 | -29.4953
19.4604 | -35.1268 | -29.5026
19.4665 | -35.1273 | -29.5094
19.4726 | -35.1277 | -29.5163
19.4787 | -35.1282 | -29.5235
19.4848 | -35.1286 | -29.5303
19.4909 | -35.1291 | -29.5372
19.497 | -35.1296 | -29.5442
19.5031 | -35.13 | -29.551
19.5092 | -35.1305 | -29.5581
19.5153 | -35.131 | -29.565
19.5214 | -35.1314 | -29.5716
19.5275 | -35.1319 | -29.5785
19.5336 | -35.1323 | -29.5856
19.5397 | -35.1328 | -29.5925
19.5458 | -35.1332 | -29.5994
19.5519 | -35.1337 | -29.6064
19.5581 | -35.1342 | -29.613
19.5642 | -35.1346 | -29.62
19.5703 | -35.1351 | -29.6268
19.5764 | -35.1355 | -29.6338
19.5825 | -35.136 | -29.6407
19.5886 | -35.1365 | -29.6476
19.5947 | -35.137 | -29.6541
19.6008 | -35.1374 | -29.661
19.6069 | -35.1379 | -29.6681
19.613 | -35.1384 | -29.675
19.6191 | -35.1388 | -29.6819
19.6252 | -35.1393 | -29.6888
19.6313 | -35.1397 | -29.6955
19.6374 | -35.1402 | -29.7026
19.6435 | -35.1406 | -29.7095
19.6496 | -35.1411 | -29.7165
19.6557 | -35.1416 | -29.7234
19.6618 | -35.142 | -29.73
19.6679 | -35.1424 | -29.737
19.674 | -35.1429 | -29.7438
19.6801 | -35.1434 | -29.7507
19.6862 | -35.1439 | -29.7577
19.6923 | -35.1443 | -29.7647
19.6984 | -35.1448 | -29.7716
19.7046 | -35.1453 | -29.7786
19.7107 | -35.1457 | -29.7854
19.7168 | -35.1462 | -29.7923
19.7229 | -35.1467 | -29.7993
19.729 | -35.1472 | -29.8061
19.7351 | -35.1476 | -29.8127
19.7412 | -35.1481 | -29.8196
19.7473 | -35.1486 | -29.8265
19.7534 | -35.149 | -29.8335
19.7595 | -35.1495 | -29.8405
19.7656 | -35.15 | -29.8475
19.7717 | -35.1504 | -29.8542
19.7778 | -35.1508 | -29.8611
19.7839 | -35.1513 | -29.8679
19.79 | -35.1518 | -29.8747
19.7961 | -35.1523 | -29.8817
19.8022 | -35.1528 | -29.8883
19.8083 | -35.1533 | -29.8952
19.8144 | -35.1537 | -29.9019
19.8205 | -35.1542 | -29.9088
19.8266 | -35.1547 | -29.9155
19.8327 | -35.1551 | -29.9223
19.8388 | -35.1555 | -29.9293
19.845 | -35.156 | -29.9361
19.8511 | -35.1564 | -29.9431
19.8572 | -35.1569 | -29.9499
19.8633 | -35.1574 | -29.9568
19.8694 | -35.1578 | -29.9638
19.8755 | -35.1583 | -29.9705
19.8816 | -35.1588 | -29.9775
19.8877 | -35.1592 | -29.9845
19.8938 | -35.1597 | -29.9914
19.8999 | -35.1602 | -29.9984
19.906 | -35.1606 | -30.0053
19.9121 | -35.1611 | -30.0121
19.9182 | -35.1616 | -30.0191
19.9243 | -35.162 | -30.026
19.9304 | -35.1624 | -30.0327
19.9365 | -35.1629 | -30.0396
19.9426 | -35.1634 | -30.0464
19.9487 | -35.1639 | -30.0532
19.9548 | -35.1643 | -30.0601
19.9609 | -35.1648 | -30.0669
19.967 | -35.1652 | -30.0738
19.9731 | -35.1657 | -30.0807
19.9792 | -35.1662 | -30.0873
19.9853 | -35.1666 | -30.094
19.9915 | -35.1671 | -30.1007
19.9976 | -35.1676 | -30.1075
20.0037 | -35.168 | -30.1144
20.0098 | -35.1685 | -30.1212
20.0159 | -35.1689 | -30.1279
20.022 | -35.1694 | -30.1347
20.0281 | -35.1698 | -30.1415
20.0342 | -35.1703 | -30.1484
20.0403 | -35.1708 | -30.1551
20.0464 | -35.1712 | -30.1621
20.0525 | -35.1717 | -30.1687
20.0586 | -35.1722 | -30.1756
20.0647 | -35.1726 | -30.1824
20.0708 | -35.1731 | -30.1895
20.0769 | -35.1735 | -30.1964
20.083 | -35.174 | -30.2027
20.0891 | -35.1744 | -30.2096
20.0952 | -35.1749 | -30.2163
20.1013 | -35.1753 | -30.2232
20.1074 | -35.1758 | -30.23
20.1135 | -35.1763 | -30.237
20.1196 | -35.1767 | -30.2435
20.1257 | -35.1772 | -30.2502
20.1319 | -35.1777 | -30.2572
20.138 | -35.1781 | -30.264
20.1441 | -35.1786 | -30.2709
20.1502 | -35.179 | -30.2775
20.1563 | -35.1795 | -30.2843
20.1624 | -35.18 | -30.291
20.1685 | -35.1804 | -30.2977
20.1746 | -35.1809 | -30.3046
20.1807 | -35.1813 | -30.3114
20.1868 | -35.1818 | -30.3182
20.1929 | -35.1822 | -30.3248
20.199 | -35.1827 | -30.3315
20.2051 | -35.1831 | -30.3381
20.2112 | -35.1836 | -30.3448
20.2173 | -35.1841 | -30.3516
20.2234 | -35.1845 | -30.3584
20.2295 | -35.185 | -30.365
20.2356 | -35.1854 | -30.372
20.2417 | -35.1859 | -30.3787
20.2478 | -35.1864 | -30.3855
20.2539 | -35.1868 | -30.3923
20.26 | -35.1873 | -30.3988
20.2661 | -35.1877 | -30.4057
20.2723 | -35.1882 | -30.4124
20.2784 | -35.1886 | -30.419
20.2845 | -35.1891 | -30.4257
20.2906 | -35.1896 | -30.4325
20.2967 | -35.19 | -30.439
20.3028 | -35.1905 | -30.4457
20.3089 | -35.191 | -30.4523
20.315 | -35.1914 | -30.4591
20.3211 | -35.1919 | -30.4657
20.3272 | -35.1923 | -30.4725
20.3333 | -35.1928 | -30.4789
20.3394 | -35.1933 | -30.4855
20.3455 | -35.1937 | -30.4924
20.3516 | -35.1942 | -30.499
20.3577 | -35.1946 | -30.5056
20.3638 | -35.1951 | -30.5124
20.3699 | -35.1955 | -30.5189
20.376 | -35.196 | -30.5257
20.3821 | -35.1965 | -30.5324
20.3882 | -35.1969 | -30.5391
20.3943 | -35.1974 | -30.5456
20.4004 | -35.1979 | -30.5519
20.4065 | -35.1983 | -30.5584
20.4126 | -35.1988 | -30.5649
20.4188 | -35.1992 | -30.5718
20.4249 | -35.1997 | -30.5784
20.431 | -35.2001 | -30.5849
20.4371 | -35.2006 | -30.591
20.4432 | -35.201 | -30.5976
20.4493 | -35.2015 | -30.6043
20.4554 | -35.2019 | -30.611
20.4615 | -35.2024 | -30.6178
20.4676 | -35.2028 | -30.6245
20.4737 | -35.2033 | -30.6308
20.4798 | -35.2037 | -30.6374
20.4859 | -35.2042 | -30.6441
20.492 | -35.2046 | -30.6508
20.4981 | -35.2051 | -30.6574
20.5042 | -35.2055 | -30.6641
20.5103 | -35.206 | -30.6706
20.5164 | -35.2064 | -30.6772
20.5225 | -35.2068 | -30.6839
20.5286 | -35.2073 | -30.6906
20.5347 | -35.2078 | -30.6972
20.5408 | -35.2082 | -30.7035
20.5469 | -35.2086 | -30.7103
20.553 | -35.2091 | -30.717
20.5592 | -35.2095 | -30.7237
20.5653 | -35.21 | -30.7303
20.5714 | -35.2105 | -30.7369
20.5775 | -35.2109 | -30.7433
20.5836 | -35.2114 | -30.75
20.5897 | -35.2119 | -30.7567
20.5958 | -35.2123 | -30.7633
20.6019 | -35.2128 | -30.77
20.608 | -35.2132 | -30.7765
20.6141 | -35.2136 | -30.7829
20.6202 | -35.2141 | -30.7896
20.6263 | -35.2145 | -30.7961
20.6324 | -35.215 | -30.8028
20.6385 | -35.2154 | -30.8094
20.6446 | -35.2159 | -30.816
20.6507 | -35.2163 | -30.8225
20.6568 | -35.2168 | -30.8289
20.6629 | -35.2172 | -30.8354
20.669 | -35.2177 | -30.842
20.6751 | -35.2181 | -30.8484
20.6812 | -35.2186 | -30.8549
20.6873 | -35.219 | -30.8615
20.6934 | -35.2195 | -30.8682
20.6995 | -35.2199 | -30.8746
20.7057 | -35.2204 | -30.8813
20.7118 | -35.2208 | -30.8881
20.7179 | -35.2213 | -30.8944
20.724 | -35.2217 | -30.9009
20.7301 | -35.2222 | -30.9075
20.7362 | -35.2226 | -30.9142
20.7423 | -35.2231 | -30.9207
20.7484 | -35.2236 | -30.9272
20.7545 | -35.224 | -30.9335
20.7606 | -35.2245 | -30.9401
20.7667 | -35.2249 | -30.9469
20.7728 | -35.2254 | -30.9534
20.7789 | -35.2259 | -30.96
20.785 | -35.2263 | -30.9665
20.7911 | -35.2268 | -30.973
20.7972 | -35.2272 | -30.9794
20.8033 | -35.2277 | -30.9861
20.8094 | -35.2281 | -30.9928
20.8155 | -35.2286 | -30.9992
20.8216 | -35.229 | -31.0057
20.8277 | -35.2294 | -31.012
20.8338 | -35.2299 | -31.0184
20.8399 | -35.2303 | -31.025
20.8461 | -35.2308 | -31.0315
20.8522 | -35.2312 | -31.0382
20.8583 | -35.2317 | -31.0443
20.8644 | -35.2321 | -31.0509
20.8705 | -35.2326 | -31.0575
20.8766 | -35.233 | -31.0639
20.8827 | -35.2335 | -31.0704
20.8888 | -35.2339 | -31.0769
20.8949 | -35.2343 | -31.0832
20.901 | -35.2347 | -31.0896
20.9071 | -35.2352 | -31.0961
20.9132 | -35.2356 | -31.1028
20.9193 | -35.2361 | -31.1093
20.9254 | -35.2365 | -31.1157
20.9315 | -35.237 | -31.1217
20.9376 | -35.2374 | -31.1281
20.9437 | -35.2379 | -31.1347
20.9498 | -35.2383 | -31.1412
20.9559 | -35.2388 | -31.1478
20.962 | -35.2392 | -31.1543
20.9681 | -35.2397 | -31.1608
20.9742 | -35.2402 | -31.1671
20.9803 | -35.2406 | -31.1736
20.9864 | -35.241 | -31.1802
20.9926 | -35.2415 | -31.187
20.9987 | -35.2419 | -31.1931
21.0048 | -35.2424 | -31.1997
21.0109 | -35.2428 | -31.2063
21.017 | -35.2433 | -31.2126
21.0231 | -35.2437 | -31.219
21.0292 | -35.2442 | -31.2255
21.0353 | -35.2446 | -31.2315
21.0414 | -35.2451 | -31.2379
21.0475 | -35.2455 | -31.2446
21.0536 | -35.246 | -31.2508
21.0597 | -35.2464 | -31.2573
21.0658 | -35.2469 | -31.2638
21.0719 | -35.2473 | -31.2698
21.078 | -35.2477 | -31.2766
21.0841 | -35.2482 | -31.2829
21.0902 | -35.2486 | -31.2895
21.0963 | -35.2491 | -31.2957
21.1024 | -35.2496 | -31.3021
21.1085 | -35.25 | -31.3083
21.1146 | -35.2505 | -31.3147
21.1207 | -35.2509 | -31.321
21.1268 | -35.2513 | -31.3275
21.133 | -35.2518 | -31.3339
21.1391 | -35.2523 | -31.34
21.1452 | -35.2527 | -31.3462
21.1513 | -35.2532 | -31.3526
21.1574 | -35.2536 | -31.359
21.1635 | -35.254 | -31.3655
21.1696 | -35.2545 | -31.3718
21.1757 | -35.2549 | -31.3779
21.1818 | -35.2553 | -31.3843
21.1879 | -35.2558 | -31.3907
21.194 | -35.2563 | -31.397
21.2001 | -35.2567 | -31.4034
21.2062 | -35.2572 | -31.4097
21.2123 | -35.2576 | -31.4154
21.2184 | -35.258 | -31.4216
21.2245 | -35.2585 | -31.428
21.2306 | -35.2589 | -31.4343
21.2367 | -35.2593 | -31.4409
21.2428 | -35.2598 | -31.447
21.2489 | -35.2602 | -31.4532
21.255 | -35.2607 | -31.4597
21.2611 | -35.2611 | -31.4658
21.2672 | -35.2615 | -31.4721
21.2733 | -35.262 | -31.4787
21.2795 | -35.2624 | -31.4846
21.2856 | -35.2628 | -31.4909
21.2917 | -35.2633 | -31.4973
21.2978 | -35.2637 | -31.5037
21.3039 | -35.2642 | -31.5098
21.31 | -35.2646 | -31.5162
21.3161 | -35.2651 | -31.5223
21.3222 | -35.2655 | -31.5287
21.3283 | -35.2659 | -31.5349
21.3344 | -35.2664 | -31.541
21.3405 | -35.2668 | -31.5475
21.3466 | -35.2672 | -31.554
21.3527 | -35.2677 | -31.56
21.3588 | -35.2681 | -31.5663
21.3649 | -35.2685 | -31.5725
21.371 | -35.269 | -31.5787
21.3771 | -35.2694 | -31.5851
21.3832 | -35.2698 | -31.5913
21.3893 | -35.2702 | -31.5975
21.3954 | -35.2706 | -31.6039
21.4015 | -35.2711 | -31.6101
21.4076 | -35.2715 | -31.6163
21.4137 | -35.272 | -31.6227
21.4199 | -35.2724 | -31.6292
21.426 | -35.2729 | -31.6349
21.4321 | -35.2733 | -31.6411
21.4382 | -35.2738 | -31.6474
21.4443 | -35.2742 | -31.6537
21.4504 | -35.2746 | -31.66
21.4565 | -35.2751 | -31.6664
21.4626 | -35.2756 | -31.6726
21.4687 | -35.276 | -31.6788
21.4748 | -35.2764 | -31.6851
21.4809 | -35.2768 | -31.6912
21.487 | -35.2773 | -31.6975
21.4931 | -35.2777 | -31.7035
21.4992 | -35.2781 | -31.7097
21.5053 | -35.2786 | -31.7161
21.5114 | -35.279 | -31.7224
21.5175 | -35.2794 | -31.7287
21.5236 | -35.2799 | -31.7348
21.5297 | -35.2803 | -31.7407
21.5358 | -35.2807 | -31.7469
21.5419 | -35.2811 | -31.7531
21.548 | -35.2815 | -31.7594
21.5541 | -35.2819 | -31.7658
21.5602 | -35.2823 | -31.7723
21.5664 | -35.2827 | -31.7781
21.5725 | -35.2832 | -31.7845
21.5786 | -35.2836 | -31.791
21.5847 | -35.284 | -31.7974
21.5908 | -35.2844 | -31.8036
21.5969 | -35.2848 | -31.8094
21.603 | -35.2853 | -31.8158
21.6091 | -35.2857 | -31.8221
21.6152 | -35.2861 | -31.8284
21.6213 | -35.2865 | -31.8348
21.6274 | -35.287 | -31.8412
21.6335 | -35.2873 | -31.8472
21.6396 | -35.2878 | -31.8535
21.6457 | -35.2882 | -31.8601
21.6518 | -35.2886 | -31.8662
21.6579 | -35.2891 | -31.8725
21.664 | -35.2895 | -31.8789
21.6701 | -35.2899 | -31.8847
21.6762 | -35.2904 | -31.8909
21.6823 | -35.2908 | -31.8975
21.6884 | -35.2913 | -31.9039
21.6945 | -35.2917 | -31.9102
21.7006 | -35.2922 | -31.9165
21.7068 | -35.2926 | -31.9225
21.7129 | -35.2931 | -31.9286
21.719 | -35.2935 | -31.9348
21.7251 | -35.2939 | -31.9414
21.7312 | -35.2943 | -31.9477
21.7373 | -35.2947 | -31.9536
21.7434 | -35.2951 | -31.9599
21.7495 | -35.2956 | -31.9661
21.7556 | -35.296 | -31.9722
21.7617 | -35.2964 | -31.9786
21.7678 | -35.2968 | -31.9846
21.7739 | -35.2973 | -31.9904
21.78 | -35.2977 | -31.9969
21.7861 | -35.2981 | -32.0031
21.7922 | -35.2986 | -32.0094
21.7983 | -35.299 | -32.0154
21.8044 | -35.2995 | -32.0216
21.8105 | -35.2999 | -32.0273
21.8166 | -35.3004 | -32.0332
21.8227 | -35.3008 | -32.0395
21.8288 | -35.3013 | -32.0457
21.8349 | -35.3017 | -32.052
21.841 | -35.3021 | -32.0582
21.8471 | -35.3025 | -32.0643
21.8533 | -35.303 | -32.0705
21.8594 | -35.3034 | -32.0767
21.8655 | -35.3038 | -32.0828
21.8716 | -35.3042 | -32.0891
21.8777 | -35.3046 | -32.0951
21.8838 | -35.305 | -32.1012
21.8899 | -35.3055 | -32.1078
21.896 | -35.3059 | -32.114
21.9021 | -35.3063 | -32.1204
21.9082 | -35.3068 | -32.1266
21.9143 | -35.3072 | -32.1322
21.9204 | -35.3076 | -32.138
21.9265 | -35.3081 | -32.1441
21.9326 | -35.3085 | -32.1502
21.9387 | -35.309 | -32.1565
21.9448 | -35.3094 | -32.1625
21.9509 | -35.3099 | -32.1684
21.957 | -35.3103 | -32.1745
21.9631 | -35.3107 | -32.1807
21.9692 | -35.3111 | -32.1869
21.9753 | -35.3116 | -32.1928
21.9814 | -35.312 | -32.1988
21.9875 | -35.3124 | -32.2046
21.9937 | -35.3128 | -32.2108
21.9998 | -35.3133 | -32.217
22.0059 | -35.3137 | -32.2232
22.012 | -35.3142 | -32.2292
22.0181 | -35.3146 | -32.2351
22.0242 | -35.315 | -32.2411
22.0303 | -35.3155 | -32.2471
22.0364 | -35.3159 | -32.2532
22.0425 | -35.3163 | -32.2591
22.0486 | -35.3168 | -32.2652
22.0547 | -35.3172 | -32.2712
22.0608 | -35.3176 | -32.2771
22.0669 | -35.318 | -32.2831
22.073 | -35.3185 | -32.2889
22.0791 | -35.3189 | -32.2949
22.0852 | -35.3193 | -32.3008
22.0913 | -35.3197 | -32.3066
22.0974 | -35.3201 | -32.3128
22.1035 | -35.3206 | -32.319
22.1096 | -35.321 | -32.3252
22.1157 | -35.3214 | -32.3313
22.1218 | -35.3218 | -32.3376
22.1279 | -35.3223 | -32.3431
22.134 | -35.3227 | -32.3491
22.1402 | -35.3231 | -32.3552
22.1463 | -35.3236 | -32.3614
22.1524 | -35.324 | -32.3676
22.1585 | -35.3244 | -32.3739
22.1646 | -35.3248 | -32.3796
22.1707 | -35.3252 | -32.3858
22.1768 | -35.3257 | -32.3921
22.1829 | -35.3261 | -32.3979
22.189 | -35.3265 | -32.4041
22.1951 | -35.3269 | -32.4096
22.2012 | -35.3274 | -32.4158
22.2073 | -35.3278 | -32.4218
22.2134 | -35.3282 | -32.4279
22.2195 | -35.3287 | -32.4338
22.2256 | -35.3291 | -32.44
22.2317 | -35.3295 | -32.446
22.2378 | -35.33 | -32.452
22.2439 | -35.3304 | -32.4581
22.25 | -35.3308 | -32.4642
22.2561 | -35.3312 | -32.47
22.2622 | -35.3317 | -32.4761
22.2683 | -35.332 | -32.4823
22.2744 | -35.3325 | -32.4884
22.2806 | -35.3329 | -32.4947
22.2867 | -35.3334 | -32.5008
22.2928 | -35.3338 | -32.5066
22.2989 | -35.3342 | -32.5125
22.305 | -35.3347 | -32.5183
22.3111 | -35.3351 | -32.5245
22.3172 | -35.3356 | -32.5306
22.3233 | -35.336 | -32.5367
22.3294 | -35.3364 | -32.5429
22.3355 | -35.3368 | -32.5486
22.3416 | -35.3372 | -32.5546
22.3477 | -35.3377 | -32.5606
22.3538 | -35.3381 | -32.5667
22.3599 | -35.3385 | -32.5725
22.366 | -35.3389 | -32.5783
22.3721 | -35.3393 | -32.5842
22.3782 | -35.3398 | -32.5903
22.3843 | -35.3402 | -32.5964
22.3904 | -35.3406 | -32.6024
22.3965 | -35.341 | -32.6086
22.4026 | -35.3414 | -32.6147
22.4087 | -35.3419 | -32.6204
22.4148 | -35.3423 | -32.6264
22.4209 | -35.3427 | -32.6325
22.4271 | -35.3431 | -32.6387
22.4332 | -35.3435 | -32.6447
22.4393 | -35.3439 | -32.6507
22.4454 | -35.3444 | -32.6566
22.4515 | -35.3448 | -32.6628
22.4576 | -35.3452 | -32.6687
22.4637 | -35.3456 | -32.6748
22.4698 | -35.346 | -32.6808
22.4759 | -35.3465 | -32.6867
22.482 | -35.3469 | -32.6927
22.4881 | -35.3473 | -32.6986
22.4942 | -35.3477 | -32.7047
22.5003 | -35.3481 | -32.7108
22.5064 | -35.3486 | -32.7167
22.5125 | -35.349 | -32.7224
22.5186 | -35.3494 | -32.7285
22.5247 | -35.3498 | -32.7345
22.5308 | -35.3502 | -32.7405
22.5369 | -35.3506 | -32.7464
22.543 | -35.3511 | -32.7524
22.5491 | -35.3515 | -32.7581
22.5552 | -35.3519 | -32.764
22.5613 | -35.3523 | -32.77
22.5675 | -35.3528 | -32.776
22.5736 | -35.3532 | -32.7819
22.5797 | -35.3536 | -32.7879
22.5858 | -35.3541 | -32.7936
22.5919 | -35.3545 | -32.7996
22.598 | -35.3549 | -32.8057
22.6041 | -35.3553 | -32.8117
22.6102 | -35.3558 | -32.8179
22.6163 | -35.3562 | -32.8236
22.6224 | -35.3566 | -32.8295
22.6285 | -35.3571 | -32.8354
22.6346 | -35.3575 | -32.8414
22.6407 | -35.3579 | -32.8474
22.6468 | -35.3583 | -32.8533
22.6529 | -35.3587 | -32.859
22.659 | -35.3591 | -32.8649
22.6651 | -35.3595 | -32.8709
22.6712 | -35.3599 | -32.8766
22.6773 | -35.3604 | -32.8826
22.6834 | -35.3608 | -32.8887
22.6895 | -35.3612 | -32.8943
22.6956 | -35.3616 | -32.9002
22.7017 | -35.362 | -32.9062
22.7079 | -35.3624 | -32.912
22.714 | -35.3629 | -32.9178
22.7201 | -35.3633 | -32.9238
22.7262 | -35.3636 | -32.9297
22.7323 | -35.3641 | -32.9357
22.7384 | -35.3645 | -32.9415
22.7445 | -35.3649 | -32.9475
22.7506 | -35.3654 | -32.9535
22.7567 | -35.3658 | -32.9594
22.7628 | -35.3662 | -32.9651
22.7689 | -35.3666 | -32.9709
22.775 | -35.367 | -32.9767
22.7811 | -35.3674 | -32.9826
22.7872 | -35.3678 | -32.9886
22.7933 | -35.3683 | -32.9941
22.7994 | -35.3687 | -32.9999
22.8055 | -35.3691 | -33.0058
22.8116 | -35.3695 | -33.0117
22.8177 | -35.3699 | -33.0176
22.8238 | -35.3703 | -33.0235
22.8299 | -35.3708 | -33.0292
22.836 | -35.3712 | -33.0351
22.8421 | -35.3716 | -33.041
22.8482 | -35.372 | -33.047
22.8544 | -35.3724 | -33.0528
22.8605 | -35.3728 | -33.0587
22.8666 | -35.3732 | -33.0645
22.8727 | -35.3737 | -33.0703
22.8788 | -35.3741 | -33.076
22.8849 | -35.3745 | -33.0818
22.891 | -35.3749 | -33.0877
22.8971 | -35.3754 | -33.0937
22.9032 | -35.3758 | -33.0994
22.9093 | -35.3762 | -33.1053
22.9154 | -35.3766 | -33.1111
22.9215 | -35.377 | -33.117
22.9276 | -35.3774 | -33.123
22.9337 | -35.3778 | -33.1286
22.9398 | -35.3782 | -33.1344
22.9459 | -35.3786 | -33.1404
22.952 | -35.3791 | -33.1462
22.9581 | -35.3795 | -33.1523
22.9642 | -35.3799 | -33.1582
22.9703 | -35.3803 | -33.1639
22.9764 | -35.3807 | -33.1697
22.9825 | -35.3811 | -33.1755
22.9886 | -35.3816 | -33.1814
22.9948 | -35.382 | -33.1872
23.0009 | -35.3824 | -33.1931
23.007 | -35.3828 | -33.1989
23.0131 | -35.3832 | -33.2047
23.0192 | -35.3836 | -33.2105
23.0253 | -35.384 | -33.2165
23.0314 | -35.3845 | -33.2223
23.0375 | -35.3849 | -33.2283
23.0436 | -35.3853 | -33.234
23.0497 | -35.3857 | -33.24
23.0558 | -35.3861 | -33.246
23.0619 | -35.3865 | -33.2519
23.068 | -35.3869 | -33.2577
23.0741 | -35.3873 | -33.2632
23.0802 | -35.3877 | -33.2693
23.0863 | -35.3881 | -33.275
23.0924 | -35.3885 | -33.2809
23.0985 | -35.3889 | -33.2868
23.1046 | -35.3893 | -33.2926
23.1107 | -35.3897 | -33.2983
23.1168 | -35.3901 | -33.3042
23.1229 | -35.3905 | -33.31
23.129 | -35.391 | -33.3159
23.1351 | -35.3914 | -33.3216
23.1413 | -35.3918 | -33.3273
23.1474 | -35.3922 | -33.3331
23.1535 | -35.3926 | -33.3389
23.1596 | -35.393 | -33.3447
23.1657 | -35.3934 | -33.3507
23.1718 | -35.3939 | -33.3566
23.1779 | -35.3943 | -33.3626
23.184 | -35.3947 | -33.3682
23.1901 | -35.3951 | -33.3741
23.1962 | -35.3955 | -33.3798
23.2023 | -35.3959 | -33.3856
23.2084 | -35.3963 | -33.3917
23.2145 | -35.3967 | -33.3973
23.2206 | -35.3971 | -33.4031
23.2267 | -35.3975 | -33.4089
23.2328 | -35.398 | -33.4146
23.2389 | -35.3984 | -33.4205
23.245 | -35.3988 | -33.4264
23.2511 | -35.3992 | -33.432
23.2572 | -35.3996 | -33.4378
23.2633 | -35.4 | -33.4437
23.2694 | -35.4004 | -33.4496
23.2755 | -35.4008 | -33.4555
23.2817 | -35.4012 | -33.4613
23.2878 | -35.4016 | -33.4671
23.2939 | -35.4021 | -33.473
23.3 | -35.4025 | -33.4789
23.3061 | -35.4029 | -33.4848
23.3122 | -35.4033 | -33.4905
23.3183 | -35.4037 | -33.4964
23.3244 | -35.4041 | -33.5022
23.3305 | -35.4045 | -33.5081
23.3366 | -35.4049 | -33.5139
23.3427 | -35.4053 | -33.5197
23.3488 | -35.4057 | -33.5256
23.3549 | -35.4062 | -33.5315
23.361 | -35.4066 | -33.5372
23.3671 | -35.407 | -33.5431
23.3732 | -35.4074 | -33.5488
23.3793 | -35.4078 | -33.5545
23.3854 | -35.4082 | -33.5604
23.3915 | -35.4086 | -33.566
23.3976 | -35.409 | -33.5719
23.4037 | -35.4095 | -33.5775
23.4098 | -35.4099 | -33.5833
23.4159 | -35.4103 | -33.5892
23.422 | -35.4107 | -33.595
23.4282 | -35.4111 | -33.6006
23.4343 | -35.4116 | -33.6064
23.4404 | -35.4119 | -33.6123
23.4465 | -35.4124 | -33.618
23.4526 | -35.4128 | -33.6238
23.4587 | -35.4132 | -33.6295
23.4648 | -35.4136 | -33.6351
23.4709 | -35.414 | -33.6409
23.477 | -35.4144 | -33.6465
23.4831 | -35.4148 | -33.6523
23.4892 | -35.4152 | -33.6581
23.4953 | -35.4156 | -33.6639
23.5014 | -35.416 | -33.6694
23.5075 | -35.4164 | -33.6751
23.5136 | -35.4169 | -33.681
23.5197 | -35.4173 | -33.6867
23.5258 | -35.4177 | -33.6926
23.5319 | -35.4181 | -33.6983
23.538 | -35.4185 | -33.7041
23.5441 | -35.4189 | -33.7101
23.5502 | -35.4192 | -33.7159
23.5563 | -35.4197 | -33.7217
23.5624 | -35.42 | -33.7276
23.5686 | -35.4205 | -33.7332
23.5747 | -35.4209 | -33.739
23.5808 | -35.4212 | -33.7448
23.5869 | -35.4216 | -33.7505
23.593 | -35.422 | -33.7562
23.5991 | -35.4225 | -33.762
23.6052 | -35.4228 | -33.7677
23.6113 | -35.4233 | -33.7735
23.6174 | -35.4236 | -33.7793
23.6235 | -35.424 | -33.785
23.6296 | -35.4244 | -33.7908
23.6357 | -35.4248 | -33.7966
23.6418 | -35.4253 | -33.8024
23.6479 | -35.4257 | -33.8081
23.654 | -35.4261 | -33.8138
23.6601 | -35.4265 | -33.8195
23.6662 | -35.4269 | -33.8252
23.6723 | -35.4273 | -33.8308
23.6784 | -35.4277 | -33.8366
23.6845 | -35.4281 | -33.8424
23.6906 | -35.4285 | -33.848
23.6967 | -35.4289 | -33.8537
23.7028 | -35.4294 | -33.8594
23.7089 | -35.4297 | -33.8649
23.7151 | -35.4301 | -33.8706
23.7212 | -35.4306 | -33.8765
23.7273 | -35.431 | -33.8822
23.7334 | -35.4314 | -33.888
23.7395 | -35.4318 | -33.8939
23.7456 | -35.4322 | -33.8995
23.7517 | -35.4326 | -33.9052
23.7578 | -35.433 | -33.911
23.7639 | -35.4334 | -33.9168
23.77 | -35.4338 | -33.9225
23.7761 | -35.4342 | -33.9283
23.7822 | -35.4346 | -33.9339
23.7883 | -35.435 | -33.9397
23.7944 | -35.4354 | -33.9455
23.8005 | -35.4358 | -33.9512
23.8066 | -35.4362 | -33.9569
23.8127 | -35.4366 | -33.9627
23.8188 | -35.437 | -33.9685
23.8249 | -35.4374 | -33.9743
23.831 | -35.4378 | -33.98
23.8371 | -35.4382 | -33.9858
23.8432 | -35.4386 | -33.9915
23.8493 | -35.439 | -33.997
23.8555 | -35.4394 | -34.0027
23.8616 | -35.4398 | -34.0085
23.8677 | -35.4402 | -34.0141
23.8738 | -35.4406 | -34.0199
23.8799 | -35.441 | -34.0259
23.886 | -35.4414 | -34.0315
23.8921 | -35.4418 | -34.0373
23.8982 | -35.4422 | -34.0429
23.9043 | -35.4426 | -34.0486
23.9104 | -35.443 | -34.0542
23.9165 | -35.4434 | -34.06
23.9226 | -35.4438 | -34.0656
23.9287 | -35.4442 | -34.0715
23.9348 | -35.4445 | -34.0772
23.9409 | -35.4449 | -34.083
23.947 | -35.4453 | -34.0887
23.9531 | -35.4457 | -34.0943
23.9592 | -35.4461 | -34.1001
23.9653 | -35.4465 | -34.1059
23.9714 | -35.4469 | -34.1115
23.9775 | -35.4473 | -34.1173
23.9836 | -35.4478 | -34.123
23.9897 | -35.4482 | -34.1287
23.9958 | -35.4486 | -34.1345
24.002 | -35.4489 | -34.1402
24.0081 | -35.4493 | -34.1459
24.0142 | -35.4497 | -34.1517
24.0203 | -35.4501 | -34.1573
24.0264 | -35.4505 | -34.163
24.0325 | -35.4509 | -34.1687
24.0386 | -35.4513 | -34.1745
24.0447 | -35.4517 | -34.1801
24.0508 | -35.4521 | -34.1858
24.0569 | -35.4525 | -34.1916
24.063 | -35.4529 | -34.1972
24.0691 | -35.4533 | -34.203
24.0752 | -35.4537 | -34.2088
24.0813 | -35.4541 | -34.2145
24.0874 | -35.4545 | -34.2202
24.0935 | -35.4549 | -34.2259
24.0996 | -35.4553 | -34.2312
24.1057 | -35.4557 | -34.2368
24.1118 | -35.4561 | -34.2426
24.1179 | -35.4565 | -34.2484
24.124 | -35.4569 | -34.2541
24.1301 | -35.4573 | -34.2597
24.1362 | -35.4577 | -34.2655
24.1424 | -35.4581 | -34.2712
24.1485 | -35.4585 | -34.2771
24.1546 | -35.4589 | -34.2828
24.1607 | -35.4593 | -34.2885
24.1668 | -35.4597 | -34.294
24.1729 | -35.4601 | -34.2996
24.179 | -35.4605 | -34.3052
24.1851 | -35.4609 | -34.3108
24.1912 | -35.4612 | -34.3165
24.1973 | -35.4616 | -34.3222
24.2034 | -35.4621 | -34.3279
24.2095 | -35.4624 | -34.3337
24.2156 | -35.4628 | -34.3395
24.2217 | -35.4632 | -34.3452
24.2278 | -35.4636 | -34.3508
24.2339 | -35.464 | -34.3564
24.24 | -35.4644 | -34.362
24.2461 | -35.4648 | -34.3678
24.2522 | -35.4652 | -34.3736
24.2583 | -35.4656 | -34.3793
24.2644 | -35.466 | -34.3851
24.2705 | -35.4664 | -34.3905
24.2766 | -35.4668 | -34.3962
24.2827 | -35.4672 | -34.402
24.2889 | -35.4676 | -34.4077
24.295 | -35.4679 | -34.4132
24.3011 | -35.4683 | -34.419
24.3072 | -35.4687 | -34.4248
24.3133 | -35.4691 | -34.4305
24.3194 | -35.4695 | -34.4361
24.3255 | -35.4699 | -34.442
24.3316 | -35.4703 | -34.4477
24.3377 | -35.4707 | -34.4533
24.3438 | -35.471 | -34.4589
24.3499 | -35.4714 | -34.4647
24.356 | -35.4718 | -34.4704
24.3621 | -35.4722 | -34.4761
24.3682 | -35.4726 | -34.4817
24.3743 | -35.473 | -34.4874
24.3804 | -35.4733 | -34.493
24.3865 | -35.4737 | -34.4987
24.3926 | -35.4741 | -34.5046
24.3987 | -35.4745 | -34.5103
24.4048 | -35.4749 | -34.516
24.4109 | -35.4753 | -34.5215
24.417 | -35.4757 | -34.5273
24.4231 | -35.4761 | -34.5328
24.4293 | -35.4765 | -34.5384
24.4354 | -35.4769 | -34.5442
24.4415 | -35.4773 | -34.5499
24.4476 | -35.4777 | -34.5554
24.4537 | -35.478 | -34.5611
24.4598 | -35.4784 | -34.5667
24.4659 | -35.4788 | -34.5725
24.472 | -35.4792 | -34.5782
24.4781 | -35.4796 | -34.5838
24.4842 | -35.4799 | -34.5894
24.4903 | -35.4803 | -34.5952
24.4964 | -35.4807 | -34.6009
24.5025 | -35.4811 | -34.6066
24.5086 | -35.4815 | -34.6122
24.5147 | -35.4819 | -34.6178
24.5208 | -35.4823 | -34.6232
24.5269 | -35.4826 | -34.6288
24.533 | -35.483 | -34.6344
24.5391 | -35.4834 | -34.6401
24.5452 | -35.4838 | -34.6459
24.5513 | -35.4842 | -34.6514
24.5574 | -35.4846 | -34.657
24.5635 | -35.485 | -34.6627
24.5696 | -35.4853 | -34.6684
24.5758 | -35.4857 | -34.6742
24.5819 | -35.4861 | -34.6798
24.588 | -35.4865 | -34.6854
24.5941 | -35.4869 | -34.6911
24.6002 | -35.4873 | -34.6967
24.6063 | -35.4877 | -34.7024
24.6124 | -35.4881 | -34.708
24.6185 | -35.4885 | -34.7136
24.6246 | -35.4889 | -34.7192
24.6307 | -35.4892 | -34.7249
24.6368 | -35.4896 | -34.7308
24.6429 | -35.49 | -34.7364
24.649 | -35.4904 | -34.7421
24.6551 | -35.4908 | -34.7476
24.6612 | -35.4912 | -34.7531
24.6673 | -35.4916 | -34.7587
24.6734 | -35.492 | -34.7644
24.6795 | -35.4924 | -34.7701
24.6856 | -35.4927 | -34.7757
24.6917 | -35.4931 | -34.7814
24.6978 | -35.4935 | -34.7868
24.7039 | -35.4939 | -34.7924
24.71 | -35.4942 | -34.7982
24.7162 | -35.4946 | -34.8039
24.7223 | -35.495 | -34.8095
24.7284 | -35.4954 | -34.815
24.7345 | -35.4958 | -34.8206
24.7406 | -35.4962 | -34.8263
24.7467 | -35.4966 | -34.8319
24.7528 | -35.497 | -34.8376
24.7589 | -35.4973 | -34.8433
24.765 | -35.4977 | -34.8488
24.7711 | -35.4982 | -34.8544
24.7772 | -35.4985 | -34.8599
24.7833 | -35.4989 | -34.8654
24.7894 | -35.4993 | -34.8711
24.7955 | -35.4997 | -34.8768
24.8016 | -35.5001 | -34.8824
24.8077 | -35.5005 | -34.888
24.8138 | -35.501 | -34.8937
24.8199 | -35.5014 | -34.8993
24.826 | -35.5017 | -34.9049
24.8321 | -35.5021 | -34.9105
24.8382 | -35.5025 | -34.9158
24.8443 | -35.5029 | -34.9214
24.8504 | -35.5033 | -34.9271
24.8565 | -35.5036 | -34.9328
24.8627 | -35.504 | -34.9384
24.8688 | -35.5044 | -34.9439
24.8749 | -35.5048 | -34.9494
24.881 | -35.5052 | -34.955
24.8871 | -35.5055 | -34.9606
24.8932 | -35.5059 | -34.9663
24.8993 | -35.5063 | -34.9718
24.9054 | -35.5067 | -34.9773
24.9115 | -35.5071 | -34.9831
24.9176 | -35.5075 | -34.9887
24.9237 | -35.5079 | -34.9944
24.9298 | -35.5082 | -35
24.9359 | -35.5086 | -35.0058
24.942 | -35.509 | -35.0114
24.9481 | -35.5094 | -35.0171
24.9542 | -35.5098 | -35.0227
24.9603 | -35.5102 | -35.0283
24.9664 | -35.5105 | -35.0337
24.9725 | -35.5109 | -35.0395
24.9786 | -35.5113 | -35.0449
24.9847 | -35.5117 | -35.0505
24.9908 | -35.512 | -35.0562
24.9969 | -35.5124 | -35.0618
25.0031 | -35.5128 | -35.0674
25.0092 | -35.5132 | -35.0729
25.0153 | -35.5136 | -35.0784
25.0214 | -35.514 | -35.084
25.0275 | -35.5143 | -35.0896
25.0336 | -35.5147 | -35.0952
25.0397 | -35.5151 | -35.1009
25.0458 | -35.5155 | -35.1063
25.0519 | -35.5159 | -35.112
25.058 | -35.5162 | -35.1178
25.0641 | -35.5166 | -35.1234
25.0702 | -35.517 | -35.1288
25.0763 | -35.5173 | -35.1345
25.0824 | -35.5177 | -35.1399
25.0885 | -35.5181 | -35.1455
25.0946 | -35.5185 | -35.1511
25.1007 | -35.5188 | -35.1568
25.1068 | -35.5192 | -35.1625
25.1129 | -35.5196 | -35.1681
25.119 | -35.5199 | -35.1735
25.1251 | -35.5203 | -35.1791
25.1312 | -35.5207 | -35.1848
25.1373 | -35.5211 | -35.1903
25.1435 | -35.5215 | -35.196
25.1496 | -35.5219 | -35.2014
25.1557 | -35.5222 | -35.207
25.1618 | -35.5226 | -35.2126
25.1679 | -35.523 | -35.2182
25.174 | -35.5234 | -35.2239
25.1801 | -35.5237 | -35.2296
25.1862 | -35.5241 | -35.2352
25.1923 | -35.5245 | -35.2407
25.1984 | -35.5249 | -35.2463
25.2045 | -35.5252 | -35.2518
25.2106 | -35.5256 | -35.2574
25.2167 | -35.526 | -35.2631
25.2228 | -35.5264 | -35.2687
25.2289 | -35.5267 | -35.2742
25.235 | -35.5272 | -35.2798
25.2411 | -35.5275 | -35.2855
25.2472 | -35.5279 | -35.291
25.2533 | -35.5283 | -35.2966
25.2594 | -35.5286 | -35.3021
25.2655 | -35.529 | -35.3077
25.2716 | -35.5294 | -35.3132
25.2777 | -35.5298 | -35.3188
25.2838 | -35.5302 | -35.3243
25.29 | -35.5305 | -35.3299
25.2961 | -35.5309 | -35.3352
25.3022 | -35.5313 | -35.3406
25.3083 | -35.5317 | -35.3462
25.3144 | -35.532 | -35.3518
25.3205 | -35.5324 | -35.3574
25.3266 | -35.5328 | -35.3629
25.3327 | -35.5332 | -35.3685
25.3388 | -35.5335 | -35.3742
25.3449 | -35.5339 | -35.3796
25.351 | -35.5343 | -35.3852
25.3571 | -35.5347 | -35.3909
25.3632 | -35.535 | -35.3964
25.3693 | -35.5354 | -35.4019
25.3754 | -35.5358 | -35.4074
25.3815 | -35.5362 | -35.413
25.3876 | -35.5365 | -35.4187
25.3937 | -35.5369 | -35.4241
25.3998 | -35.5373 | -35.4295
25.4059 | -35.5376 | -35.4351
25.412 | -35.538 | -35.4406
25.4181 | -35.5384 | -35.4461
25.4242 | -35.5388 | -35.4516
25.4304 | -35.5392 | -35.4571
25.4365 | -35.5395 | -35.4625
25.4426 | -35.5399 | -35.468
25.4487 | -35.5403 | -35.4737
25.4548 | -35.5406 | -35.4792
25.4609 | -35.541 | -35.4848
25.467 | -35.5414 | -35.4902
25.4731 | -35.5418 | -35.4957
25.4792 | -35.5421 | -35.5013
25.4853 | -35.5425 | -35.5069
25.4914 | -35.5429 | -35.5125
25.4975 | -35.5433 | -35.5179
25.5036 | -35.5436 | -35.5234
25.5097 | -35.544 | -35.529
25.5158 | -35.5444 | -35.5346
25.5219 | -35.5448 | -35.5403
25.528 | -35.5451 | -35.5459
25.5341 | -35.5455 | -35.5515
25.5402 | -35.5459 | -35.557
25.5463 | -35.5463 | -35.5624
25.5524 | -35.5467 | -35.5679
25.5585 | -35.547 | -35.5735
25.5646 | -35.5474 | -35.5791
25.5707 | -35.5477 | -35.5846
25.5769 | -35.5481 | -35.5901
25.583 | -35.5485 | -35.5955
25.5891 | -35.5488 | -35.601
25.5952 | -35.5492 | -35.6066
25.6013 | -35.5496 | -35.6122
25.6074 | -35.55 | -35.6175
25.6135 | -35.5503 | -35.623
25.6196 | -35.5507 | -35.6286
25.6257 | -35.5511 | -35.634
25.6318 | -35.5515 | -35.6395
25.6379 | -35.5518 | -35.6451
25.644 | -35.5522 | -35.6505
25.6501 | -35.5525 | -35.656
25.6562 | -35.5529 | -35.6615
25.6623 | -35.5533 | -35.667
25.6684 | -35.5537 | -35.6725
25.6745 | -35.5541 | -35.6781
25.6806 | -35.5544 | -35.6834
25.6867 | -35.5548 | -35.6888
25.6928 | -35.5552 | -35.6944
25.6989 | -35.5555 | -35.6998
25.705 | -35.5559 | -35.7054
25.7111 | -35.5562 | -35.7109
25.7173 | -35.5566 | -35.7164
25.7234 | -35.557 | -35.722
25.7295 | -35.5574 | -35.7276
25.7356 | -35.5577 | -35.7331
25.7417 | -35.5581 | -35.7385
25.7478 | -35.5584 | -35.7439
25.7539 | -35.5588 | -35.7493
25.76 | -35.5592 | -35.7547
25.7661 | -35.5595 | -35.7603
25.7722 | -35.5599 | -35.7658
25.7783 | -35.5602 | -35.7713
25.7844 | -35.5606 | -35.7769
25.7905 | -35.561 | -35.7825
25.7966 | -35.5613 | -35.788
25.8027 | -35.5617 | -35.7936
25.8088 | -35.562 | -35.7992
25.8149 | -35.5624 | -35.8047
25.821 | -35.5627 | -35.8101
25.8271 | -35.5631 | -35.8157
25.8332 | -35.5635 | -35.8212
25.8393 | -35.5638 | -35.8267
25.8454 | -35.5642 | -35.8322
25.8515 | -35.5646 | -35.8377
25.8576 | -35.5649 | -35.843
25.8638 | -35.5653 | -35.8486
25.8699 | -35.5656 | -35.854
25.876 | -35.566 | -35.8594
25.8821 | -35.5664 | -35.8649
25.8882 | -35.5667 | -35.8701
25.8943 | -35.5671 | -35.8757
25.9004 | -35.5674 | -35.8813
25.9065 | -35.5678 | -35.8869
25.9126 | -35.5682 | -35.8923
25.9187 | -35.5685 | -35.8977
25.9248 | -35.5689 | -35.903
25.9309 | -35.5693 | -35.9084
25.937 | -35.5696 | -35.9139
25.9431 | -35.57 | -35.9194
25.9492 | -35.5704 | -35.9249
25.9553 | -35.5707 | -35.9305
25.9614 | -35.571 | -35.9357
25.9675 | -35.5714 | -35.9411
25.9736 | -35.5717 | -35.9467
25.9797 | -35.5721 | -35.9522
25.9858 | -35.5724 | -35.9576
25.9919 | -35.5728 | -35.9631
25.998 | -35.5731 | -35.9685
26.0042 | -35.5735 | -35.974
26.0103 | -35.5739 | -35.9795
26.0164 | -35.5742 | -35.9849
26.0225 | -35.5746 | -35.9903
26.0286 | -35.5749 | -35.9956
26.0347 | -35.5753 | -36.0008
26.0408 | -35.5756 | -36.0062
26.0469 | -35.576 | -36.0117
26.053 | -35.5764 | -36.017
26.0591 | -35.5768 | -36.0224
26.0652 | -35.5771 | -36.0276
26.0713 | -35.5775 | -36.033
26.0774 | -35.5778 | -36.0384
26.0835 | -35.5782 | -36.0439
26.0896 | -35.5785 | -36.0493
26.0957 | -35.5789 | -36.0547
26.1018 | -35.5792 | -36.0599
26.1079 | -35.5796 | -36.0654
26.114 | -35.5799 | -36.071
26.1201 | -35.5803 | -36.0764
26.1262 | -35.5806 | -36.0819
26.1323 | -35.581 | -36.0874
26.1384 | -35.5814 | -36.0927
26.1445 | -35.5817 | -36.0982
26.1507 | -35.5821 | -36.1037
26.1568 | -35.5824 | -36.1092
26.1629 | -35.5828 | -36.1145
26.169 | -35.5832 | -36.12
26.1751 | -35.5835 | -36.1253
26.1812 | -35.5839 | -36.1307
26.1873 | -35.5842 | -36.1362
26.1934 | -35.5846 | -36.1416
26.1995 | -35.5849 | -36.1471
26.2056 | -35.5853 | -36.1525
26.2117 | -35.5856 | -36.158
26.2178 | -35.586 | -36.1633
26.2239 | -35.5863 | -36.1688
26.23 | -35.5867 | -36.1742
26.2361 | -35.5871 | -36.1797
26.2422 | -35.5874 | -36.185
26.2483 | -35.5878 | -36.1902
26.2544 | -35.5881 | -36.1957
26.2605 | -35.5885 | -36.2011
26.2666 | -35.5888 | -36.2066
26.2727 | -35.5892 | -36.212
26.2788 | -35.5895 | -36.2173
26.2849 | -35.5899 | -36.2227
26.2911 | -35.5903 | -36.2281
26.2972 | -35.5906 | -36.2335
26.3033 | -35.591 | -36.2389
26.3094 | -35.5913 | -36.2443
26.3155 | -35.5917 | -36.2497
26.3216 | -35.5921 | -36.2552
26.3277 | -35.5924 | -36.2606
26.3338 | -35.5927 | -36.2659
26.3399 | -35.5931 | -36.2712
26.346 | -35.5935 | -36.2765
26.3521 | -35.5938 | -36.282
26.3582 | -35.5942 | -36.2875
26.3643 | -35.5945 | -36.2928
26.3704 | -35.5949 | -36.2982
26.3765 | -35.5952 | -36.3036
26.3826 | -35.5956 | -36.3089
26.3887 | -35.5959 | -36.3142
26.3948 | -35.5963 | -36.3196
26.4009 | -35.5966 | -36.3251
26.407 | -35.597 | -36.3304
26.4131 | -35.5974 | -36.3358
26.4192 | -35.5977 | -36.341
26.4253 | -35.5981 | -36.3464
26.4314 | -35.5984 | -36.3517
26.4376 | -35.5988 | -36.3572
26.4437 | -35.5991 | -36.3625
26.4498 | -35.5995 | -36.368
26.4559 | -35.5998 | -36.3732
26.462 | -35.6002 | -36.3786
26.4681 | -35.6005 | -36.384
26.4742 | -35.6009 | -36.3894
26.4803 | -35.6012 | -36.3948
26.4864 | -35.6015 | -36.4
26.4925 | -35.6019 | -36.4053
26.4986 | -35.6022 | -36.4106
26.5047 | -35.6026 | -36.4159
26.5108 | -35.6029 | -36.4213
26.5169 | -35.6033 | -36.4267
26.523 | -35.6036 | -36.432
26.5291 | -35.604 | -36.4374
26.5352 | -35.6043 | -36.4427
26.5413 | -35.6047 | -36.4482
26.5474 | -35.605 | -36.4536
26.5535 | -35.6054 | -36.459
26.5596 | -35.6057 | -36.4644
26.5657 | -35.606 | -36.4698
26.5718 | -35.6064 | -36.4753
26.578 | -35.6067 | -36.4806
26.5841 | -35.6071 | -36.4861
26.5902 | -35.6074 | -36.4914
26.5963 | -35.6078 | -36.4966
26.6024 | -35.6081 | -36.5019
26.6085 | -35.6085 | -36.5072
26.6146 | -35.6088 | -36.5126
26.6207 | -35.6092 | -36.5181
26.6268 | -35.6095 | -36.5235
26.6329 | -35.6098 | -36.5286
26.639 | -35.6101 | -36.534
26.6451 | -35.6105 | -36.5394
26.6512 | -35.6108 | -36.5446
26.6573 | -35.6112 | -36.55
26.6634 | -35.6115 | -36.5552
26.6695 | -35.6119 | -36.5606
26.6756 | -35.6122 | -36.5662
26.6817 | -35.6125 | -36.5716
26.6878 | -35.6129 | -36.5769
26.6939 | -35.6132 | -36.5822
26.7 | -35.6135 | -36.5874
26.7061 | -35.6139 | -36.5928
26.7122 | -35.6142 | -36.598
26.7183 | -35.6146 | -36.6033
26.7245 | -35.6149 | -36.6087
26.7306 | -35.6153 | -36.6139
26.7367 | -35.6156 | -36.619
26.7428 | -35.6159 | -36.6244
26.7489 | -35.6163 | -36.6298
26.755 | -35.6166 | -36.635
26.7611 | -35.617 | -36.6404
26.7672 | -35.6173 | -36.6458
26.7733 | -35.6176 | -36.6511
26.7794 | -35.618 | -36.6565
26.7855 | -35.6184 | -36.6618
26.7916 | -35.6187 | -36.667
26.7977 | -35.619 | -36.6722
26.8038 | -35.6194 | -36.6774
26.8099 | -35.6197 | -36.6828
26.816 | -35.62 | -36.6881
26.8221 | -35.6204 | -36.6934
26.8282 | -35.6207 | -36.6986
26.8343 | -35.621 | -36.7038
26.8404 | -35.6213 | -36.709
26.8465 | -35.6217 | -36.7142
26.8526 | -35.622 | -36.7195
26.8587 | -35.6224 | -36.7249
26.8649 | -35.6227 | -36.7303
26.871 | -35.6231 | -36.7358
26.8771 | -35.6234 | -36.7408
26.8832 | -35.6237 | -36.7461
26.8893 | -35.624 | -36.7514
26.8954 | -35.6244 | -36.7567
26.9015 | -35.6247 | -36.7622
26.9076 | -35.6251 | -36.7676
26.9137 | -35.6254 | -36.7727
26.9198 | -35.6257 | -36.778
26.9259 | -35.6261 | -36.7833
26.932 | -35.6264 | -36.7887
26.9381 | -35.6268 | -36.7941
26.9442 | -35.6271 | -36.7994
26.9503 | -35.6274 | -36.8046
26.9564 | -35.6277 | -36.8099
26.9625 | -35.6281 | -36.8153
26.9686 | -35.6284 | -36.8206
26.9747 | -35.6287 | -36.8259
26.9808 | -35.629 | -36.8311
26.9869 | -35.6294 | -36.8363
26.993 | -35.6297 | -36.8416
26.9991 | -35.63 | -36.8468
27.0052 | -35.6304 | -36.852
27.0114 | -35.6307 | -36.8573
27.0175 | -35.631 | -36.8625
27.0236 | -35.6314 | -36.8678
27.0297 | -35.6317 | -36.8731
27.0358 | -35.632 | -36.8784
27.0419 | -35.6324 | -36.8837
27.048 | -35.6327 | -36.889
27.0541 | -35.6331 | -36.894
27.0602 | -35.6334 | -36.8993
27.0663 | -35.6337 | -36.9046
27.0724 | -35.6341 | -36.9099
27.0785 | -35.6344 | -36.9153
27.0846 | -35.6347 | -36.9204
27.0907 | -35.6351 | -36.9258
27.0968 | -35.6354 | -36.9309
27.1029 | -35.6357 | -36.9362
27.109 | -35.6361 | -36.9415
27.1151 | -35.6364 | -36.9469
27.1212 | -35.6367 | -36.9519
27.1273 | -35.6371 | -36.9571
27.1334 | -35.6374 | -36.9624
27.1395 | -35.6377 | -36.9677
27.1456 | -35.6381 | -36.973
27.1518 | -35.6384 | -36.9783
27.1579 | -35.6387 | -36.9834
27.164 | -35.639 | -36.9887
27.1701 | -35.6394 | -36.9941
27.1762 | -35.6397 | -36.9993
27.1823 | -35.64 | -37.0045
27.1884 | -35.6404 | -37.0099
27.1945 | -35.6407 | -37.015
27.2006 | -35.641 | -37.0202
27.2067 | -35.6413 | -37.0255
27.2128 | -35.6417 | -37.0308
27.2189 | -35.642 | -37.036
27.225 | -35.6423 | -37.0413
27.2311 | -35.6426 | -37.0464
27.2372 | -35.643 | -37.0517
27.2433 | -35.6433 | -37.0568
27.2494 | -35.6437 | -37.0621
27.2555 | -35.644 | -37.0673
27.2616 | -35.6443 | -37.0724
27.2677 | -35.6446 | -37.0776
27.2738 | -35.6449 | -37.083
27.2799 | -35.6453 | -37.0885
27.286 | -35.6456 | -37.0937
27.2921 | -35.6459 | -37.099
27.2983 | -35.6462 | -37.104
27.3044 | -35.6466 | -37.1094
27.3105 | -35.6469 | -37.1146
27.3166 | -35.6472 | -37.1198
27.3227 | -35.6476 | -37.1249
27.3288 | -35.6479 | -37.13
27.3349 | -35.6482 | -37.135
27.341 | -35.6485 | -37.1402
27.3471 | -35.6488 | -37.1454
27.3532 | -35.6492 | -37.1506
27.3593 | -35.6495 | -37.156
27.3654 | -35.6499 | -37.1612
27.3715 | -35.6502 | -37.1661
27.3776 | -35.6505 | -37.1714
27.3837 | -35.6508 | -37.1766
27.3898 | -35.6511 | -37.1818
27.3959 | -35.6515 | -37.1871
27.402 | -35.6518 | -37.1922
27.4081 | -35.6521 | -37.1974
27.4142 | -35.6524 | -37.2025
27.4203 | -35.6528 | -37.2077
27.4264 | -35.6531 | -37.213
27.4325 | -35.6534 | -37.2182
27.4387 | -35.6537 | -37.2232
27.4448 | -35.654 | -37.2285
27.4509 | -35.6543 | -37.2337
27.457 | -35.6546 | -37.2389
27.4631 | -35.6549 | -37.244
27.4692 | -35.6552 | -37.2493
27.4753 | -35.6555 | -37.2544
27.4814 | -35.6558 | -37.2594
27.4875 | -35.6562 | -37.2648
27.4936 | -35.6565 | -37.27
27.4997 | -35.6568 | -37.2753
27.5058 | -35.6572 | -37.2805
27.5119 | -35.6575 | -37.2856
27.518 | -35.6578 | -37.2908
27.5241 | -35.6581 | -37.296
27.5302 | -35.6584 | -37.3013
27.5363 | -35.6587 | -37.3063
27.5424 | -35.659 | -37.3113
27.5485 | -35.6593 | -37.3166
27.5546 | -35.6596 | -37.3216
27.5607 | -35.66 | -37.3268
27.5668 | -35.6603 | -37.332
27.5729 | -35.6606 | -37.3374
27.5791 | -35.6609 | -37.3425
27.5852 | -35.6612 | -37.3478
27.5913 | -35.6616 | -37.353
27.5974 | -35.6619 | -37.3582
27.6035 | -35.6622 | -37.3634
27.6096 | -35.6625 | -37.3685
27.6157 | -35.6629 | -37.3735
27.6218 | -35.6632 | -37.3788
27.6279 | -35.6635 | -37.384
27.634 | -35.6638 | -37.3894
27.6401 | -35.6642 | -37.3946
27.6462 | -35.6645 | -37.3997
27.6523 | -35.6648 | -37.4046
27.6584 | -35.6651 | -37.4098
27.6645 | -35.6655 | -37.415
27.6706 | -35.6658 | -37.4201
27.6767 | -35.6661 | -37.4254
27.6828 | -35.6664 | -37.4304
27.6889 | -35.6667 | -37.4355
27.695 | -35.667 | -37.4405
27.7011 | -35.6673 | -37.4457
27.7072 | -35.6677 | -37.451
27.7133 | -35.668 | -37.4561
27.7194 | -35.6683 | -37.4612
27.7256 | -35.6686 | -37.4664
27.7317 | -35.6689 | -37.4716
27.7378 | -35.6692 | -37.4767
27.7439 | -35.6695 | -37.482
27.75 | -35.6699 | -37.487
27.7561 | -35.6702 | -37.492
27.7622 | -35.6705 | -37.4971
27.7683 | -35.6708 | -37.5023
27.7744 | -35.6711 | -37.5075
27.7805 | -35.6715 | -37.5127
27.7866 | -35.6718 | -37.5179
27.7927 | -35.6721 | -37.5227
27.7988 | -35.6724 | -37.5279
27.8049 | -35.6727 | -37.5331
27.811 | -35.673 | -37.5383
27.8171 | -35.6733 | -37.5435
27.8232 | -35.6736 | -37.5485
27.8293 | -35.6739 | -37.5536
27.8354 | -35.6743 | -37.5589
27.8415 | -35.6746 | -37.564
27.8476 | -35.6749 | -37.5692
27.8537 | -35.6752 | -37.5743
27.8598 | -35.6755 | -37.5792
27.866 | -35.6758 | -37.5844
27.8721 | -35.6761 | -37.5896
27.8782 | -35.6764 | -37.5948
27.8843 | -35.6767 | -37.5999
27.8904 | -35.6771 | -37.605
27.8965 | -35.6774 | -37.61
27.9026 | -35.6777 | -37.615
27.9087 | -35.678 | -37.6201
27.9148 | -35.6783 | -37.6251
27.9209 | -35.6786 | -37.6303
27.927 | -35.679 | -37.6354
27.9331 | -35.6792 | -37.6403
27.9392 | -35.6795 | -37.6454
27.9453 | -35.6799 | -37.6505
27.9514 | -35.6801 | -37.6556
27.9575 | -35.6804 | -37.661
27.9636 | -35.6807 | -37.666
27.9697 | -35.681 | -37.671
27.9758 | -35.6813 | -37.6761
27.9819 | -35.6816 | -37.6812
27.988 | -35.6819 | -37.6863
27.9941 | -35.6822 | -37.6915
28.0002 | -35.6825 | -37.6964
28.0063 | -35.6828 | -37.7015
28.0125 | -35.6831 | -37.7066
28.0186 | -35.6834 | -37.7117
28.0247 | -35.6837 | -37.7168
28.0308 | -35.684 | -37.722
28.0369 | -35.6843 | -37.7268
28.043 | -35.6846 | -37.7319
28.0491 | -35.6849 | -37.7371
28.0552 | -35.6852 | -37.7422
28.0613 | -35.6855 | -37.7472
28.0674 | -35.6858 | -37.7523
28.0735 | -35.6861 | -37.7572

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4874 | -42.499
21.4908 | -35.555 | -34.6677
16.4948 | -34.9412 | -27.3798
12.6602 | -34.5823 | -21.1853
9.7171 | -34.3604 | -16.3348
7.4582 | -34.2136 | -12.3936
5.7244 | -34.1058 | -9.1998
4.3936 | -34.0389 | -6.6281
3.3722 | -34.0244 | -4.597
2.5883 | -34.0154 | -3.0592
1.9866 | -34.0192 | -1.8296
1.5248 | -34.0181 | -0.84474
1.1703 | -34.0236 | -0.2019
1 | -34.0236 | 0.11384
0.89822 | -34.0086 | 0.38487
0.68941 | -33.9994 | 0.9611
0.52919 | -34.0004 | 1.4047
0.40612 | -34.0011 | 2.0115
0.31171 | -33.9883 | 2.4675
0.23928 | -34.0095 | 3.1843
0.18358 | -34.0061 | 3.6029
0.14091 | -34.0284 | 4.4758
0.10819 | -34.0312 | 5.6443
0.083017 | -34.0693 | 7.1126
0.063705 | -34.1068 | 8.9845
0.048923 | -34.181 | 11.4699
0.037479 | -34.2803 | 14.6442
0.028801 | -34.4594 | 18.6627
0.022078 | -34.7702 | 24.0075
0.017023 | -35.2488 | 30.4398
0.013018 | -36.0563 | 38.1548
0.010014 | -37.1679 | 46.0545

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003704 | -13.2759
0.007407 | -28.8775
0.011111 | -37.2722
0.014815 | -41.2525
0.018519 | -43.7111
0.022222 | -45.4126
0.025926 | -46.7184
0.02963 | -48.1116
0.033333 | -49.8438
0.037037 | -50.3256
0.040741 | -51.7681
0.044444 | -53.2056
0.048148 | -55.3276
0.051852 | -55.9973
0.055556 | -56.0425
0.059259 | -58.0765
0.062963 | -58.5632
0.066667 | -58.7354
0.07037 | -59.0677
0.074074 | -59.8684
0.077778 | -59.6493
0.081481 | -61.2283
0.085185 | -62.4986
0.088889 | -62.6365
0.092593 | -63.1246
0.096296 | -62.616
0.1 | -63.585
0.1037 | -62.9132
0.10741 | -63.1909
0.11111 | -64.3448
0.11482 | -64.8528
0.11852 | -65.3708
0.12222 | -65.5764
0.12593 | -65.8857
0.12963 | -66.5084
0.13333 | -66.4543
0.13704 | -67.7137
0.14074 | -66.9386
0.14444 | -65.812
0.14815 | -66.7895
0.15185 | -67.9995
0.15556 | -68.3942
0.15926 | -69.3175
0.16296 | -68.7318
0.16667 | -69.0744
0.17037 | -69.4898
0.17407 | -69.9974
0.17778 | -69.717
0.18148 | -69.1256
0.18518 | -69.3481
0.18889 | -70.8679
0.19259 | -70.3696
0.1963 | -69.5979
0.2 | -70.6163
0.2037 | -71.3422
0.20741 | -71.5297
0.21111 | -71.2243
0.21482 | -71.1181
0.21852 | -70.657
0.22222 | -71.8532
0.22593 | -72.8153
0.22963 | -72.2555
0.23333 | -71.7529
0.23704 | -71.9659
0.24074 | -73.2687
0.24444 | -74.1447
0.24815 | -74.3197
0.25185 | -74.1452
0.25556 | -73.8436
0.25926 | -74.5168
0.26296 | -74.3812
0.26667 | -74.208
0.27037 | -74.0919
0.27407 | -74.3931
0.27778 | -74.5947
0.28148 | -74.414
0.28519 | -73.9495
0.28889 | -73.3495
0.29259 | -73.5075
0.2963 | -73.8234
0.3 | -73.8576
0.3037 | -75.3529
0.30741 | -76.0617
0.31111 | -76.4493
0.31482 | -75.5743
0.31852 | -74.8672
0.32222 | -74.3398
0.32593 | -75.8081
0.32963 | -76.8874
0.33333 | -76.6168
0.33704 | -76.9571
0.34074 | -76.4846
0.34444 | -76.5842
0.34815 | -76.696
0.35185 | -77.1639
0.35556 | -76.5988
0.35926 | -75.5322
0.36296 | -76.0426
0.36667 | -77.2937
0.37037 | -78.1532
0.37407 | -77.0857
0.37778 | -76.9383
0.38148 | -76.9207
0.38518 | -77.5578
0.38889 | -78.1038
0.39259 | -77.8615
0.3963 | -77.0368
0.4 | -77.1076
0.4037 | -78.6344
0.40741 | -78.4679
0.41111 | -78.1593
0.41481 | -77.6966
0.41852 | -78.3169
0.42222 | -78.8863
0.42593 | -78.4694
0.42963 | -77.8404
0.43333 | -77.5849
0.43704 | -77.8716
0.44074 | -78.2482
0.44444 | -78.5467
0.44815 | -78.6001
0.45185 | -78.9342
0.45556 | -79.3506
0.45926 | -78.8375
0.46296 | -79.2052
0.46667 | -80.0287
0.47037 | -79.6473
0.47407 | -79.0397
0.47778 | -80.2407
0.48148 | -80.3006
0.48518 | -79.6797
0.48889 | -80.0683
0.49259 | -79.8173
0.4963 | -80.4241
0.5 | -80.2442
0.5037 | -79.9543
0.50741 | -79.5501
0.51111 | -79.4917
0.51482 | -79.8044
0.51852 | -80.2805
0.52222 | -79.7025
0.52593 | -79.6624
0.52963 | -80.7537
0.53333 | -80.9087
0.53704 | -80.8513
0.54074 | -80.8182
0.54444 | -81.2629
0.54815 | -81.5262
0.55185 | -81.2101
0.55556 | -80.821
0.55926 | -80.8538
0.56296 | -81.0547
0.56667 | -81.081
0.57037 | -81.3654
0.57407 | -82.0569
0.57778 | -82.5918
0.58148 | -82.4635
0.58518 | -81.8065
0.58889 | -81.6879
0.59259 | -81.8359
0.5963 | -81.4731
0.6 | -81.3839
0.6037 | -81.586
0.60741 | -81.6121
0.61111 | -81.9604
0.61482 | -82.2854
0.61852 | -82.2383
0.62222 | -82.1569
0.62593 | -82.4685
0.62963 | -82.848
0.63333 | -82.7641
0.63704 | -82.4605
0.64074 | -82.7187
0.64444 | -82.9902
0.64815 | -82.7491
0.65185 | -82.4839
0.65556 | -82.2623
0.65926 | -82.2237
0.66296 | -82.4435
0.66667 | -82.9254
0.67037 | -83.5083
0.67407 | -83.5231
0.67778 | -83.4597
0.68148 | -83.5374
0.68519 | -83.2701
0.68889 | -82.7054
0.69259 | -82.5049
0.6963 | -82.8876
0.7 | -83.174
0.7037 | -82.7226
0.70741 | -82.4113
0.71111 | -82.6531
0.71481 | -83.349
0.71852 | -83.726
0.72222 | -83.5174
0.72593 | -83.3105
0.72963 | -83.5732
0.73333 | -83.8367
0.73704 | -83.5926
0.74074 | -83.1238
0.74444 | -82.9337
0.74815 | -83.345
0.75185 | -83.7764
0.75556 | -83.67
0.75926 | -83.4933
0.76296 | -83.8024
0.76667 | -84.4087
0.77037 | -84.6477
0.77407 | -84.5773
0.77778 | -84.6936
0.78148 | -84.8077
0.78519 | -84.8039
0.78889 | -84.3226
0.79259 | -83.7826
0.7963 | -83.8462
0.8 | -84.0396
0.8037 | -83.9068
0.80741 | -83.7468
0.81111 | -83.7524
0.81481 | -83.9389
0.81852 | -84.7107
0.82222 | -85.2935
0.82593 | -85.0883
0.82963 | -85.4455
0.83333 | -86.2697
0.83704 | -86.116
0.84074 | -85.1113
0.84444 | -84.4907
0.84815 | -84.4587
0.85185 | -84.4735
0.85556 | -84.345
0.85926 | -84.0865
0.86296 | -84.0283
0.86667 | -84.7002
0.87037 | -85.5692
0.87407 | -85.5607
0.87778 | -84.8647
0.88148 | -84.7042
0.88518 | -85.0248
0.88889 | -84.9981
0.89259 | -84.7455
0.8963 | -84.7007
0.9 | -85.0503
0.9037 | -85.519
0.90741 | -85.2988
0.91111 | -85.0601
0.91482 | -85.2028
0.91852 | -85.5393
0.92222 | -85.9598
0.92593 | -86.6033
0.92963 | -87.2306
0.93333 | -87.4909
0.93704 | -87.2828
0.94074 | -86.7612
0.94444 | -86.4571
0.94815 | -86.1943
0.95185 | -85.9304
0.95556 | -86.1638
0.95926 | -86.5178
0.96296 | -86.4475
0.96667 | -86.2437
0.97037 | -85.9704
0.97407 | -85.9824
0.97778 | -86.3351
0.98148 | -86.424
0.98518 | -86.3761
0.98889 | -86.5549
0.99259 | -86.9459
0.9963 | -86.875
1 | -86.9547
1.0037 | -87.4317
1.0074 | -87.9831
1.0111 | -88.0906
1.0148 | -87.8124
1.0185 | -87.4428
1.0222 | -87.3116
1.0259 | -87.4364
1.0296 | -87.6994
1.0333 | -88.0249
1.037 | -88.2751
1.0407 | -88.2801
1.0444 | -88.0734
1.0481 | -87.6508
1.0519 | -87.182
1.0556 | -87.0998
1.0593 | -87.4411
1.063 | -87.4891
1.0667 | -87.4728
1.0704 | -87.5131
1.0741 | -87.5645
1.0778 | -87.6864
1.0815 | -87.9048
1.0852 | -88.0002
1.0889 | -87.7205
1.0926 | -87.2374
1.0963 | -87.0048
1.1 | -87.2879
1.1037 | -87.9206
1.1074 | -88.3318
1.1111 | -88.208
1.1148 | -87.8294
1.1185 | -87.5996
1.1222 | -87.7158
1.1259 | -87.9268
1.1296 | -87.9346
1.1333 | -88.0227
1.137 | -88.4838
1.1407 | -89.0466
1.1444 | -89.1106
1.1481 | -88.6985
1.1519 | -88.4575
1.1556 | -88.5748
1.1593 | -88.6536
1.163 | -88.4086
1.1667 | -88.0752
1.1704 | -87.9807
1.1741 | -88.2018
1.1778 | -88.5313
1.1815 | -88.7417
1.1852 | -88.8143
1.1889 | -88.6848
1.1926 | -88.4085
1.1963 | -88.34
1.2 | -88.5249
1.2037 | -88.578
1.2074 | -88.3933
1.2111 | -88.2233
1.2148 | -88.1463
1.2185 | -88.2021
1.2222 | -88.4285
1.2259 | -88.7166
1.2296 | -88.9952
1.2333 | -89.1658
1.237 | -89.0559
1.2407 | -88.7417
1.2444 | -88.5192
1.2481 | -88.5098
1.2519 | -88.6104
1.2556 | -88.7609
1.2593 | -88.9531
1.263 | -89.1354
1.2667 | -89.2031
1.2704 | -89.1409
1.2741 | -89.132
1.2778 | -89.2762
1.2815 | -89.502
1.2852 | -89.718
1.2889 | -89.8514
1.2926 | -89.8518
1.2963 | -89.6657
1.3 | -89.3443
1.3037 | -89.11
1.3074 | -89.1488
1.3111 | -89.4373
1.3148 | -89.734
1.3185 | -89.7311
1.3222 | -89.402
1.3259 | -89.0545
1.3296 | -88.9849
1.3333 | -89.3414
1.337 | -89.9628
1.3407 | -90.4056
1.3444 | -90.542
1.3481 | -90.6126
1.3519 | -90.7042
1.3556 | -90.6945
1.3593 | -90.4213
1.363 | -89.9783
1.3667 | -89.6985
1.3704 | -89.7918
1.3741 | -90.1321
1.3778 | -90.358
1.3815 | -90.3099
1.3852 | -90.0946
1.3889 | -89.7998
1.3926 | -89.5647
1.3963 | -89.6351
1.4 | -90.0697
1.4037 | -90.5974
1.4074 | -90.8424
1.4111 | -90.6739
1.4148 | -90.3265
1.4185 | -90.0609
1.4222 | -89.9354
1.4259 | -89.9452
1.4296 | -90.0359
1.4333 | -90.0416
1.437 | -89.7993
1.4407 | -89.5087
1.4444 | -89.5515
1.4481 | -89.9344
1.4519 | -90.3586
1.4556 | -90.5718
1.4593 | -90.4995
1.463 | -90.419
1.4667 | -90.6147
1.4704 | -91.0163
1.4741 | -91.362
1.4778 | -91.399
1.4815 | -91.0799
1.4852 | -90.7001
1.4889 | -90.481
1.4926 | -90.4589
1.4963 | -90.5555
1.5 | -90.6063
1.5037 | -90.5899
1.5074 | -90.6359
1.5111 | -90.8004
1.5148 | -91.0506
1.5185 | -91.2683
1.5222 | -91.3266
1.5259 | -91.2827
1.5296 | -91.2289
1.5333 | -91.2078
1.537 | -91.2619
1.5407 | -91.2363
1.5444 | -90.9904
1.5481 | -90.7068
1.5519 | -90.4597
1.5556 | -90.2224
1.5593 | -90.1779
1.563 | -90.445
1.5667 | -90.7899
1.5704 | -90.944
1.5741 | -90.9072
1.5778 | -90.8804
1.5815 | -91.0474
1.5852 | -91.3356
1.5889 | -91.4578
1.5926 | -91.2195
1.5963 | -90.781
1.6 | -90.7472
1.6037 | -90.8065
1.6074 | -90.9925
1.6111 | -91.208
1.6148 | -91.4254
1.6185 | -91.6342
1.6222 | -91.7963
1.6259 | -91.9207
1.6296 | -92.0152
1.6333 | -92.0209
1.637 | -91.9053
1.6407 | -91.7524
1.6444 | -91.6823
1.6481 | -91.7593
1.6519 | -91.9287
1.6556 | -92.0597
1.6593 | -92.0499
1.663 | -91.9535
1.6667 | -91.9081
1.6704 | -91.9556
1.6741 | -92.0373
1.6778 | -92.0968
1.6815 | -92.1132
1.6852 | -92.083
1.6889 | -92.0228
1.6926 | -91.9963
1.6963 | -92.0403
1.7 | -92.1001
1.7037 | -92.0949
1.7074 | -91.9613
1.7111 | -91.7493
1.7148 | -91.5355
1.7185 | -91.3873
1.7222 | -91.3229
1.7259 | -91.3511
1.7296 | -91.4838
1.7333 | -91.7112
1.737 | -91.9629
1.7407 | -92.172
1.7444 | -92.3084
1.7481 | -92.3964
1.7519 | -92.4186
1.7556 | -92.3408
1.7593 | -92.1798
1.763 | -92.0286
1.7667 | -91.9776
1.7704 | -92.0358
1.7741 | -92.1192
1.7778 | -92.172
1.7815 | -92.2567
1.7852 | -92.4598
1.7889 | -92.7469
1.7926 | -92.9707
1.7963 | -93.0511
1.8 | -93.0476
1.8037 | -93.0407
1.8074 | -93.0572
1.8111 | -93.0659
1.8148 | -93.0897
1.8185 | -93.2446
1.8222 | -93.5687
1.8259 | -93.9164
1.8296 | -94.0546
1.8333 | -93.9606
1.837 | -93.7614
1.8407 | -93.5291
1.8444 | -93.2088
1.8481 | -92.7794
1.8519 | -92.3872
1.8556 | -92.1999
1.8593 | -92.2257
1.863 | -92.3852
1.8667 | -92.5626
1.8704 | -92.6515
1.8741 | -92.6091
1.8778 | -92.502
1.8815 | -92.4325
1.8852 | -92.4539
1.8889 | -92.5518
1.8926 | -92.6552
1.8963 | -92.6709
1.9 | -92.6015
1.9037 | -92.5451
1.9074 | -92.6085
1.9111 | -92.8435
1.9148 | -93.1937
1.9185 | -93.5306
1.9222 | -93.6983
1.9259 | -93.6068
1.9296 | -93.3663
1.9333 | -93.2126
1.937 | -93.3186
1.9407 | -93.6158
1.9444 | -93.8356
1.9481 | -93.7575
1.9519 | -93.4378
1.9556 | -93.1071
1.9593 | -92.8996
1.963 | -92.849
1.9667 | -92.9721
1.9704 | -93.2097
1.9741 | -93.4096
1.9778 | -93.4475
1.9815 | -93.3558
1.9852 | -93.2339
1.9889 | -93.144
1.9926 | -93.1367
1.9963 | -93.2215
2 | -93.3476
2.0037 | -93.456
2.0074 | -93.5041
2.0111 | -93.4744
2.0148 | -93.3881
2.0185 | -93.3358
2.0222 | -93.3696
2.0259 | -93.4459
2.0296 | -93.4942
2.0333 | -93.5306
2.037 | -93.6149
2.0407 | -93.7417
2.0444 | -93.8673
2.0481 | -93.9868
2.0519 | -94.043
2.0556 | -93.9944
2.0593 | -93.8624
2.063 | -93.7349
2.0667 | -93.7104
2.0704 | -93.846
2.0741 | -94.0746
2.0778 | -94.1967
2.0815 | -94.04
2.0852 | -93.6696
2.0889 | -93.3313
2.0926 | -93.1979
2.0963 | -93.2249
2.1 | -93.2538
2.1037 | -93.2017
2.1074 | -93.1326
2.1111 | -93.1077
2.1148 | -93.0972
2.1185 | -93.1086
2.1222 | -93.1949
2.1259 | -93.3672
2.1296 | -93.5673
2.1333 | -93.7104
2.137 | -93.7145
2.1407 | -93.7194
2.1444 | -93.693
2.1481 | -93.6609
2.1519 | -93.671
2.1556 | -93.7462
2.1593 | -93.8825
2.163 | -94.0412
2.1667 | -94.1501
2.1704 | -94.1882
2.1741 | -94.1492
2.1778 | -94.057
2.1815 | -93.9238
2.1852 | -93.8109
2.1889 | -93.7986
2.1926 | -93.9054
2.1963 | -94.0825
2.2 | -94.2328
2.2037 | -94.2541
2.2074 | -94.136
2.2111 | -93.9676
2.2148 | -93.8616
2.2185 | -93.856
2.2222 | -93.9345
2.2259 | -94.0202
2.2296 | -94.0321
2.2333 | -93.9109
2.237 | -93.6808
2.2407 | -93.4657
2.2444 | -93.3846
2.2481 | -93.4653
2.2519 | -93.6485
2.2556 | -93.84
2.2593 | -93.956
2.263 | -93.9424
2.2667 | -93.8177
2.2704 | -93.6611
2.2741 | -93.5589
2.2778 | -93.5314
2.2815 | -93.5758
2.2852 | -93.6563
2.2889 | -93.7477
2.2926 | -93.8537
2.2963 | -93.9919
2.3 | -94.1613
2.3037 | -94.3405
2.3074 | -94.4844
2.3111 | -94.5771
2.3148 | -94.6194
2.3185 | -94.6299
2.3222 | -94.5879
2.3259 | -94.5055
2.3296 | -94.3993
2.3333 | -94.2974
2.337 | -94.2059
2.3407 | -94.1226
2.3444 | -94.0346
2.3481 | -93.9636
2.3519 | -93.9399
2.3556 | -94.0093
2.3593 | -94.1643
2.363 | -94.3805
2.3667 | -94.6052
2.3704 | -94.7916
2.3741 | -94.878
2.3778 | -94.8162
2.3815 | -94.5937
2.3852 | -94.2865
2.3889 | -93.9936
2.3926 | -93.8196
2.3963 | -93.813
2.4 | -93.9821
2.4037 | -94.2843
2.4074 | -94.6661
2.4111 | -95.0457
2.4148 | -95.3411
2.4185 | -95.467
2.4222 | -95.4191
2.4259 | -95.2571
2.4296 | -95.0659
2.4333 | -94.8639
2.437 | -94.6703
2.4407 | -94.5369
2.4444 | -94.4934
2.4481 | -94.5085
2.4519 | -94.5128
2.4556 | -94.4577
2.4593 | -94.3674
2.463 | -94.3136
2.4667 | -94.3762
2.4704 | -94.5716
2.4741 | -94.8658
2.4778 | -95.1606
2.4815 | -95.3276
2.4852 | -95.3173
2.4889 | -95.1866
2.4926 | -95.035
2.4963 | -94.949
2.5 | -94.9664
2.5037 | -95.0725
2.5074 | -95.2409
2.5111 | -95.46
2.5148 | -95.6909
2.5185 | -95.8739
2.5222 | -95.9367
2.5259 | -95.8646
2.5296 | -95.724
2.5333 | -95.6016
2.537 | -95.5632
2.5407 | -95.6217
2.5444 | -95.7438
2.5481 | -95.8514
2.5519 | -95.8551
2.5556 | -95.7178
2.5593 | -95.4906
2.563 | -95.2698
2.5667 | -95.113
2.5704 | -95.0172
2.5741 | -94.9364
2.5778 | -94.8306
2.5815 | -94.7089
2.5852 | -94.5905
2.5889 | -94.4955
2.5926 | -94.4662
2.5963 | -94.5282
2.6 | -94.6759
2.6037 | -94.889
2.6074 | -95.1329
2.6111 | -95.3872
2.6148 | -95.6275
2.6185 | -95.7893
2.6222 | -95.8303
2.6259 | -95.7467
2.6296 | -95.5955
2.6333 | -95.4223
2.637 | -95.2699
2.6407 | -95.1753
2.6444 | -95.1616
2.6481 | -95.2279
2.6519 | -95.3358
2.6556 | -95.44
2.6593 | -95.5241
2.663 | -95.5917
2.6667 | -95.6579
2.6704 | -95.719
2.6741 | -95.7725
2.6778 | -95.8229
2.6815 | -95.8907
2.6852 | -95.9734
2.6889 | -96.0629
2.6926 | -96.1558
2.6963 | -96.221
2.7 | -96.2554
2.7037 | -96.2482
2.7074 | -96.1959
2.7111 | -96.1041
2.7148 | -95.9675
2.7185 | -95.8275
2.7222 | -95.703
2.7259 | -95.6174
2.7296 | -95.5812
2.7333 | -95.5963
2.737 | -95.6368
2.7407 | -95.6865
2.7444 | -95.7409
2.7481 | -95.7933
2.7519 | -95.8356
2.7556 | -95.8618
2.7593 | -95.8698
2.763 | -95.8603
2.7667 | -95.8104
2.7704 | -95.7005
2.7741 | -95.5441
2.7778 | -95.3676
2.7815 | -95.2173
2.7852 | -95.1359
2.7889 | -95.1393
2.7926 | -95.2234
2.7963 | -95.3632
2.8 | -95.5254
2.8037 | -95.6819
2.8074 | -95.8052
2.8111 | -95.8841
2.8148 | -95.9251
2.8185 | -95.9391
2.8222 | -95.9379
2.8259 | -95.9339
2.8296 | -95.9208
2.8333 | -95.8902
2.837 | -95.8381
2.8407 | -95.7688
2.8444 | -95.7066
2.8481 | -95.6741
2.8519 | -95.6713
2.8556 | -95.7043
2.8593 | -95.745
2.863 | -95.7636
2.8667 | -95.7575
2.8704 | -95.7543
2.8741 | -95.7765
2.8778 | -95.8385
2.8815 | -95.9097
2.8852 | -95.9637
2.8889 | -95.9671
2.8926 | -95.9123
2.8963 | -95.8281
2.9 | -95.736
2.9037 | -95.6637
2.9074 | -95.6358
2.9111 | -95.6597
2.9148 | -95.7139
2.9185 | -95.787
2.9222 | -95.889
2.9259 | -96.0051
2.9296 | -96.1466
2.9333 | -96.293
2.937 | -96.4372
2.9407 | -96.5528
2.9444 | -96.6174
2.9481 | -96.6406
2.9519 | -96.6148
2.9556 | -96.5684
2.9593 | -96.5178
2.963 | -96.482
2.9667 | -96.455
2.9704 | -96.4381
2.9741 | -96.4498
2.9778 | -96.5024
2.9815 | -96.6162
2.9852 | -96.7697
2.9889 | -96.9273
2.9926 | -97.0635
2.9963 | -97.1483
3 | -97.1958
3.0037 | -97.2049
3.0074 | -97.1816
3.0111 | -97.1271
3.0148 | -97.0419
3.0185 | -96.9367
3.0222 | -96.8307
3.0259 | -96.7283
3.0296 | -96.65
3.0333 | -96.6281
3.037 | -96.6638
3.0407 | -96.7434
3.0444 | -96.8441
3.0481 | -96.9381
3.0519 | -97.01
3.0556 | -97.0524
3.0593 | -97.0591
3.063 | -97.0052
3.0667 | -96.8871
3.0704 | -96.738
3.0741 | -96.6033
3.0778 | -96.5161
3.0815 | -96.4875
3.0852 | -96.5094
3.0889 | -96.5611
3.0926 | -96.6201
3.0963 | -96.6795
3.1 | -96.7419
3.1037 | -96.8199
3.1074 | -96.9122
3.1111 | -97.0179
3.1148 | -97.1076
3.1185 | -97.1499
3.1222 | -97.1503
3.1259 | -97.1119
3.1296 | -97.0425
3.1333 | -96.9404
3.137 | -96.8194
3.1407 | -96.7041
3.1444 | -96.617
3.1481 | -96.5726
3.1519 | -96.5643
3.1556 | -96.591
3.1593 | -96.6582
3.163 | -96.7731
3.1667 | -96.9457
3.1704 | -97.1597
3.1741 | -97.4125
3.1778 | -97.6858
3.1815 | -97.9386
3.1852 | -98.124
3.1889 | -98.2026
3.1926 | -98.1634
3.1963 | -98.0406
3.2 | -97.862
3.2037 | -97.6352
3.2074 | -97.4536
3.2111 | -97.2912
3.2148 | -97.1668
3.2185 | -97.0919
3.2222 | -97.0705
3.2259 | -97.0829
3.2296 | -97.1223
3.2333 | -97.1848
3.237 | -97.2638
3.2407 | -97.3565
3.2444 | -97.4451
3.2481 | -97.4982
3.2519 | -97.4934
3.2556 | -97.4162
3.2593 | -97.2832
3.263 | -97.1514
3.2667 | -97.0671
3.2704 | -97.0614
3.2741 | -97.1361
3.2778 | -97.2564
3.2815 | -97.3922
3.2852 | -97.5012
3.2889 | -97.5547
3.2926 | -97.5525
3.2963 | -97.5127
3.3 | -97.4469
3.3037 | -97.374
3.3074 | -97.3291
3.3111 | -97.3317
3.3148 | -97.3719
3.3185 | -97.4293
3.3222 | -97.4906
3.3259 | -97.5495
3.3296 | -97.5911
3.3333 | -97.645
3.337 | -97.7105
3.3407 | -97.7855
3.3444 | -97.8621
3.3481 | -97.9226
3.3519 | -97.964
3.3556 | -97.9823
3.3593 | -97.9823
3.363 | -97.9516
3.3667 | -97.8997
3.3704 | -97.8471
3.3741 | -97.7947
3.3778 | -97.7527
3.3815 | -97.7191
3.3852 | -97.7078
3.3889 | -97.7013
3.3926 | -97.7133
3.3963 | -97.7634
3.4 | -97.8452
3.4037 | -97.9417
3.4074 | -98.0188
3.4111 | -98.0677
3.4148 | -98.0649
3.4185 | -98.0132
3.4222 | -97.9315
3.4259 | -97.854
3.4296 | -97.8073
3.4333 | -97.7878
3.437 | -97.792
3.4407 | -97.7971
3.4444 | -97.7874
3.4481 | -97.7426
3.4519 | -97.6636
3.4556 | -97.5676
3.4593 | -97.4772
3.463 | -97.4286
3.4667 | -97.4351
3.4704 | -97.5185
3.4741 | -97.6812
3.4778 | -97.8751
3.4815 | -98.0549
3.4852 | -98.1941
3.4889 | -98.2662
3.4926 | -98.2758
3.4963 | -98.2467
3.5 | -98.1991
3.5037 | -98.1359
3.5074 | -98.0659
3.5111 | -97.9927
3.5148 | -97.9175
3.5185 | -97.8286
3.5222 | -97.7377
3.5259 | -97.6362
3.5296 | -97.5304
3.5333 | -97.4507
3.537 | -97.4009
3.5407 | -97.3897
3.5444 | -97.4256
3.5481 | -97.5033
3.5519 | -97.6119
3.5556 | -97.7434
3.5593 | -97.8993
3.563 | -98.0508
3.5667 | -98.198
3.5704 | -98.333
3.5741 | -98.4437
3.5778 | -98.5163
3.5815 | -98.5745
3.5852 | -98.6014
3.5889 | -98.5847
3.5926 | -98.5298
3.5963 | -98.4404
3.6 | -98.3375
3.6037 | -98.2309
3.6074 | -98.1393
3.6111 | -98.0729
3.6148 | -98.0371
3.6185 | -98.025
3.6222 | -98.0359
3.6259 | -98.0637
3.6296 | -98.1136
3.6333 | -98.1819
3.637 | -98.2634
3.6407 | -98.3773
3.6444 | -98.5073
3.6481 | -98.6442
3.6519 | -98.7771
3.6556 | -98.8973
3.6593 | -98.9939
3.663 | -99.0497
3.6667 | -99.0693
3.6704 | -99.0484
3.6741 | -99.0061
3.6778 | -98.9368
3.6815 | -98.8606
3.6852 | -98.7883
3.6889 | -98.7316
3.6926 | -98.6879
3.6963 | -98.6304
3.7 | -98.5551
3.7037 | -98.4702
3.7074 | -98.3842
3.7111 | -98.3212
3.7148 | -98.3049
3.7185 | -98.3311
3.7222 | -98.3713
3.7259 | -98.42
3.7296 | -98.4637
3.7333 | -98.4947
3.737 | -98.4941
3.7407 | -98.5091
3.7444 | -98.518
3.7481 | -98.5118
3.7519 | -98.5003
3.7556 | -98.4907
3.7593 | -98.4834
3.763 | -98.4785
3.7667 | -98.4891
3.7704 | -98.5145
3.7741 | -98.5658
3.7778 | -98.6445
3.7815 | -98.7465
3.7852 | -98.8543
3.7889 | -98.9538
3.7926 | -99.0477
3.7963 | -99.1206
3.8 | -99.1738
3.8037 | -99.2099
3.8074 | -99.2341
3.8111 | -99.2497
3.8148 | -99.2528
3.8185 | -99.2581
3.8222 | -99.2635
3.8259 | -99.2598
3.8296 | -99.2481
3.8333 | -99.2195
3.837 | -99.1633
3.8407 | -99.0846
3.8444 | -99.0004
3.8481 | -98.9131
3.8519 | -98.8253
3.8556 | -98.7408
3.8593 | -98.6771
3.863 | -98.6318
3.8667 | -98.612
3.8704 | -98.6136
3.8741 | -98.6471
3.8778 | -98.7167
3.8815 | -98.8121
3.8852 | -98.9316
3.8889 | -99.0708
3.8926 | -99.2154
3.8963 | -99.3598
3.9 | -99.4796
3.9037 | -99.5691
3.9074 | -99.6148
3.9111 | -99.6154
3.9148 | -99.5662
3.9185 | -99.4651
3.9222 | -99.3274
3.9259 | -99.1657
3.9296 | -99.0083
3.9333 | -98.876
3.937 | -98.7769
3.9407 | -98.7119
3.9444 | -98.6913
3.9481 | -98.7021
3.9519 | -98.732
3.9556 | -98.7747
3.9593 | -98.8163
3.963 | -98.8569
3.9667 | -98.8822
3.9704 | -98.9077
3.9741 | -98.9395
3.9778 | -98.9755
3.9815 | -99.0053
3.9852 | -99.02
3.9889 | -99.026
3.9926 | -99.0079
3.9963 | -98.9667
4 | -98.9143
4.0037 | -98.8657
4.0074 | -98.819
4.0111 | -98.7983
4.0148 | -98.8162
4.0185 | -98.8737
4.0222 | -98.9737
4.0259 | -99.1102
4.0296 | -99.2653
4.0333 | -99.4015
4.037 | -99.5131
4.0407 | -99.5855
4.0444 | -99.6071
4.0481 | -99.5619
4.0519 | -99.4543
4.0556 | -99.3199
4.0593 | -99.1745
4.063 | -99.0282
4.0667 | -98.911
4.0704 | -98.8405
4.0741 | -98.8098
4.0778 | -98.8157
4.0815 | -98.863
4.0852 | -98.9451
4.0889 | -99.0454
4.0926 | -99.1477
4.0963 | -99.2421
4.1 | -99.3166
4.1037 | -99.3651
4.1074 | -99.3913
4.1111 | -99.4001
4.1148 | -99.3919
4.1185 | -99.3638
4.1222 | -99.3252
4.1259 | -99.2842
4.1296 | -99.2383
4.1333 | -99.1991
4.137 | -99.1767
4.1407 | -99.1818
4.1444 | -99.2017
4.1481 | -99.2254
4.1519 | -99.2614
4.1556 | -99.302
4.1593 | -99.3411
4.163 | -99.385
4.1667 | -99.4462
4.1704 | -99.5179
4.1741 | -99.6012
4.1778 | -99.682
4.1815 | -99.7375
4.1852 | -99.7646
4.1889 | -99.7583
4.1926 | -99.7188
4.1963 | -99.6591
4.2 | -99.598
4.2037 | -99.5471
4.2074 | -99.5204
4.2111 | -99.5266
4.2148 | -99.5455
4.2185 | -99.571
4.2222 | -99.5837
4.2259 | -99.5733
4.2296 | -99.5405
4.2333 | -99.4944
4.237 | -99.4393
4.2407 | -99.3808
4.2444 | -99.3335
4.2481 | -99.2988
4.2519 | -99.2719
4.2556 | -99.2492
4.2593 | -99.2236
4.263 | -99.1879
4.2667 | -99.1513
4.2704 | -99.1094
4.2741 | -99.0797
4.2778 | -99.0567
4.2815 | -99.0416
4.2852 | -99.0394
4.2889 | -99.0461
4.2926 | -99.0532
4.2963 | -99.0538
4.3 | -99.054
4.3037 | -99.0522
4.3074 | -99.054
4.3111 | -99.0581
4.3148 | -99.0642
4.3185 | -99.0738
4.3222 | -99.0934
4.3259 | -99.1214
4.3296 | -99.1686
4.3333 | -99.2336
4.337 | -99.3192
4.3407 | -99.4258
4.3444 | -99.5416
4.3481 | -99.6601
4.3519 | -99.7817
4.3556 | -99.8932
4.3593 | -99.9872
4.363 | -100.0595
4.3667 | -100.1094
4.3704 | -100.1217
4.3741 | -100.0995
4.3778 | -100.0524
4.3815 | -99.9849
4.3852 | -99.909
4.3889 | -99.8498
4.3926 | -99.8081
4.3963 | -99.7947
4.4 | -99.7994
4.4037 | -99.8052
4.4074 | -99.8129
4.4111 | -99.8086
4.4148 | -99.7914
4.4185 | -99.7522
4.4222 | -99.6872
4.4259 | -99.605
4.4296 | -99.5169
4.4333 | -99.4314
4.437 | -99.3437
4.4407 | -99.2713
4.4444 | -99.2215
4.4481 | -99.19
4.4519 | -99.1846
4.4556 | -99.2141
4.4593 | -99.2766
4.463 | -99.3733
4.4667 | -99.5014
4.4704 | -99.6546
4.4741 | -99.8334
4.4778 | -100.0079
4.4815 | -100.1631
4.4852 | -100.2864
4.4889 | -100.3694
4.4926 | -100.413
4.4963 | -100.4185
4.5 | -100.3944
4.5037 | -100.3469
4.5074 | -100.2999
4.5111 | -100.2481
4.5148 | -100.1959
4.5185 | -100.1517
4.5222 | -100.1168
4.5259 | -100.0837
4.5296 | -100.0499
4.5333 | -100.0236
4.537 | -99.9951
4.5407 | -99.9652
4.5444 | -99.9263
4.5481 | -99.8891
4.5519 | -99.8572
4.5556 | -99.8314
4.5593 | -99.8148
4.563 | -99.8058
4.5667 | -99.7991
4.5704 | -99.8102
4.5741 | -99.8411
4.5778 | -99.8904
4.5815 | -99.9601
4.5852 | -100.0543
4.5889 | -100.1574
4.5926 | -100.2621
4.5963 | -100.3537
4.6 | -100.4288
4.6037 | -100.477
4.6074 | -100.5053
4.6111 | -100.5136
4.6148 | -100.4993
4.6185 | -100.4677
4.6222 | -100.4239
4.6259 | -100.3718
4.6296 | -100.316
4.6333 | -100.2551
4.637 | -100.2023
4.6407 | -100.1492
4.6444 | -100.1003
4.6481 | -100.0566
4.6519 | -100.0105
4.6556 | -99.9599
4.6593 | -99.9193
4.663 | -99.8839
4.6667 | -99.8593
4.6704 | -99.8492
4.6741 | -99.8434
4.6778 | -99.8456
4.6815 | -99.8552
4.6852 | -99.8677
4.6889 | -99.8856
4.6926 | -99.9039
4.6963 | -99.9254
4.7 | -99.9495
4.7037 | -99.9721
4.7074 | -99.9893
4.7111 | -99.9977
4.7148 | -100.0046
4.7185 | -100.0151
4.7222 | -100.0323
4.7259 | -100.0622
4.7296 | -100.098
4.7333 | -100.146
4.737 | -100.1941
4.7407 | -100.2397
4.7444 | -100.2776
4.7481 | -100.2967
4.7519 | -100.3027
4.7556 | -100.3064
4.7593 | -100.2964
4.763 | -100.2847
4.7667 | -100.2657
4.7704 | -100.2417
4.7741 | -100.2162
4.7778 | -100.1947
4.7815 | -100.1933
4.7852 | -100.2159
4.7889 | -100.2654
4.7926 | -100.344
4.7963 | -100.4399
4.8 | -100.548
4.8037 | -100.6532
4.8074 | -100.7317
4.8111 | -100.8181
4.8148 | -100.8839
4.8185 | -100.9301
4.8222 | -100.9561
4.8259 | -100.9618
4.8296 | -100.9515
4.8333 | -100.9321
4.837 | -100.9017
4.8407 | -100.873
4.8444 | -100.8452
4.8481 | -100.8272
4.8519 | -100.8221
4.8556 | -100.8219
4.8593 | -100.8255
4.863 | -100.8387
4.8667 | -100.871
4.8704 | -100.9168
4.8741 | -100.9651
4.8778 | -101.0008
4.8815 | -101.0178
4.8852 | -101.0215
4.8889 | -101.0046
4.8926 | -100.974
4.8963 | -100.9386
4.9 | -100.9023
4.9037 | -100.8695
4.9074 | -100.8467
4.9111 | -100.8385
4.9148 | -100.8386
4.9185 | -100.8394
4.9222 | -100.848
4.9259 | -100.8553
4.9296 | -100.8582
4.9333 | -100.8622
4.937 | -100.8566
4.9407 | -100.8355
4.9444 | -100.7992
4.9481 | -100.7459
4.9519 | -100.6852
4.9556 | -100.6268
4.9593 | -100.5708
4.963 | -100.52
4.9667 | -100.489
4.9704 | -100.4834
4.9741 | -100.4936
4.9778 | -100.5177
4.9815 | -100.5595
4.9852 | -100.6073
4.9889 | -100.6567
4.9926 | -100.7018
4.9963 | -100.7372
5 | -100.7554
5.0037 | -100.7676
5.0074 | -100.7726
5.0111 | -100.7751
5.0148 | -100.7881
5.0185 | -100.8096
5.0222 | -100.825
5.0259 | -100.8506
5.0296 | -100.8688
5.0333 | -100.8745
5.037 | -100.8688
5.0407 | -100.8536
5.0444 | -100.8341
5.0481 | -100.8117
5.0519 | -100.8006
5.0556 | -100.7944
5.0593 | -100.7847
5.063 | -100.7789
5.0667 | -100.7751
5.0704 | -100.773
5.0741 | -100.7673
5.0778 | -100.7597
5.0815 | -100.7534
5.0852 | -100.7485
5.0889 | -100.7356
5.0926 | -100.7117
5.0963 | -100.6836
5.1 | -100.6564
5.1037 | -100.6262
5.1074 | -100.6125
5.1111 | -100.6167
5.1148 | -100.6372
5.1185 | -100.6765
5.1222 | -100.7318
5.1259 | -100.8008
5.1296 | -100.8662
5.1333 | -100.9288
5.137 | -100.983
5.1407 | -101.0215
5.1444 | -101.0453
5.1481 | -101.055
5.1519 | -101.0584
5.1556 | -101.0556
5.1593 | -101.0425
5.163 | -101.0322
5.1667 | -101.0262
5.1704 | -101.0157
5.1741 | -100.9964
5.1778 | -100.9759
5.1815 | -100.9478
5.1852 | -100.9168
5.1889 | -100.8833
5.1926 | -100.8502
5.1963 | -100.8298
5.2 | -100.8215
5.2037 | -100.832
5.2074 | -100.8684
5.2111 | -100.9213
5.2148 | -100.9868
5.2185 | -101.0539
5.2222 | -101.132
5.2259 | -101.2113
5.2296 | -101.2734
5.2333 | -101.3169
5.237 | -101.3393
5.2407 | -101.341
5.2444 | -101.3218
5.2481 | -101.2967
5.2519 | -101.2663
5.2556 | -101.2328
5.2593 | -101.2109
5.263 | -101.1882
5.2667 | -101.1722
5.2704 | -101.166
5.2741 | -101.157
5.2778 | -101.1507
5.2815 | -101.1461
5.2852 | -101.1445
5.2889 | -101.1532
5.2926 | -101.1681
5.2963 | -101.1883
5.3 | -101.1976
5.3037 | -101.208
5.3074 | -101.2201
5.3111 | -101.232
5.3148 | -101.2436
5.3185 | -101.2493
5.3222 | -101.2603
5.3259 | -101.2599
5.3296 | -101.2515
5.3333 | -101.2397
5.337 | -101.2157
5.3407 | -101.1718
5.3444 | -101.1441
5.3481 | -101.1146
5.3519 | -101.0812
5.3556 | -101.0471
5.3593 | -101.0221
5.363 | -101.009
5.3667 | -101.0071
5.3704 | -101.0082
5.3741 | -101.012
5.3778 | -101.024
5.3815 | -101.0369
5.3852 | -101.0425
5.3889 | -101.0491
5.3926 | -101.0548
5.3963 | -101.064
5.4 | -101.0751
5.4037 | -101.0823
5.4074 | -101.0929
5.4111 | -101.1036
5.4148 | -101.1077
5.4185 | -101.1117
5.4222 | -101.1217
5.4259 | -101.139
5.4296 | -101.1585
5.4333 | -101.1926
5.437 | -101.2433
5.4407 | -101.3046
5.4444 | -101.365
5.4481 | -101.4304
5.4519 | -101.4964
5.4556 | -101.5596
5.4593 | -101.6144
5.463 | -101.652
5.4667 | -101.6805
5.4704 | -101.6972
5.4741 | -101.7015
5.4778 | -101.6983
5.4815 | -101.6833
5.4852 | -101.6619
5.4889 | -101.6415
5.4926 | -101.6185
5.4963 | -101.5913
5.5 | -101.5641
5.5037 | -101.5342
5.5074 | -101.5018
5.5111 | -101.4658
5.5148 | -101.4348
5.5185 | -101.4114
5.5222 | -101.39
5.5259 | -101.386
5.5296 | -101.3971
5.5333 | -101.4118
5.537 | -101.4346
5.5407 | -101.4663
5.5444 | -101.5025
5.5481 | -101.5366
5.5519 | -101.5582
5.5556 | -101.5678
5.5593 | -101.5569
5.563 | -101.5296
5.5667 | -101.499
5.5704 | -101.466
5.5741 | -101.4298
5.5778 | -101.3979
5.5815 | -101.3783
5.5852 | -101.3711
5.5889 | -101.3772
5.5926 | -101.3985
5.5963 | -101.4347
5.6 | -101.4754
5.6037 | -101.5107
5.6074 | -101.5477
5.6111 | -101.5671
5.6148 | -101.5815
5.6185 | -101.5905
5.6222 | -101.5931
5.6259 | -101.59
5.6296 | -101.5865
5.6333 | -101.5836
5.637 | -101.5943
5.6407 | -101.6084
5.6444 | -101.6206
5.6481 | -101.6246
5.6519 | -101.6186
5.6556 | -101.6059
5.6593 | -101.5885
5.663 | -101.5678
5.6667 | -101.5365
5.6704 | -101.5028
5.6741 | -101.4699
5.6778 | -101.4394
5.6815 | -101.4217
5.6852 | -101.4161
5.6889 | -101.4277
5.6926 | -101.4565
5.6963 | -101.5067
5.7 | -101.5684
5.7037 | -101.6459
5.7074 | -101.7368
5.7111 | -101.8424
5.7148 | -101.9492
5.7185 | -102.0478
5.7222 | -102.1271
5.7259 | -102.188
5.7296 | -102.2226
5.7333 | -102.2222
5.737 | -102.1907
5.7407 | -102.1293
5.7444 | -102.0416
5.7481 | -101.9441
5.7519 | -101.8518
5.7556 | -101.7743
5.7593 | -101.716
5.763 | -101.6751
5.7667 | -101.6508
5.7704 | -101.6516
5.7741 | -101.6735
5.7778 | -101.7106
5.7815 | -101.7596
5.7852 | -101.8123
5.7889 | -101.871
5.7926 | -101.9301
5.7963 | -101.9841
5.8 | -102.0353
5.8037 | -102.0821
5.8074 | -102.1119
5.8111 | -102.1217
5.8148 | -102.1147
5.8185 | -102.0953
5.8222 | -102.0541
5.8259 | -101.9939
5.8296 | -101.927
5.8333 | -101.8492
5.837 | -101.7698
5.8407 | -101.6963
5.8444 | -101.6338
5.8481 | -101.5813
5.8519 | -101.5453
5.8556 | -101.5216
5.8593 | -101.507
5.863 | -101.5029
5.8667 | -101.5063
5.8704 | -101.5213
5.8741 | -101.5545
5.8778 | -101.5799
5.8815 | -101.6117
5.8852 | -101.6512
5.8889 | -101.7001
5.8926 | -101.757
5.8963 | -101.8216
5.9 | -101.8849
5.9037 | -101.9443
5.9074 | -101.9968
5.9111 | -102.0455
5.9148 | -102.0869
5.9185 | -102.1136
5.9222 | -102.1321
5.9259 | -102.1414
5.9296 | -102.1373
5.9333 | -102.119
5.937 | -102.1036
5.9407 | -102.0857
5.9444 | -102.0654
5.9481 | -102.0538
5.9519 | -102.0452
5.9556 | -102.0406
5.9593 | -102.0385
5.963 | -102.0411
5.9667 | -102.0473
5.9704 | -102.0539
5.9741 | -102.0653
5.9778 | -102.0753
5.9815 | -102.0799
5.9852 | -102.0842
5.9889 | -102.091
5.9926 | -102.1002
5.9963 | -102.1007
6 | -102.0996
6.0037 | -102.0975
6.0074 | -102.0948
6.0111 | -102.0894
6.0148 | -102.0775
6.0185 | -102.0641
6.0222 | -102.0575
6.0259 | -102.0594
6.0296 | -102.057
6.0333 | -102.0593
6.037 | -102.0662
6.0407 | -102.0691
6.0444 | -102.0751
6.0481 | -102.0752
6.0519 | -102.0753
6.0556 | -102.0814
6.0593 | -102.0839
6.063 | -102.0834
6.0667 | -102.0779
6.0704 | -102.0761
6.0741 | -102.073
6.0778 | -102.0656
6.0815 | -102.0568
6.0852 | -102.0467
6.0889 | -102.0367
6.0926 | -102.0192
6.0963 | -102.0019
6.1 | -101.9945
6.1037 | -101.995
6.1074 | -102.0096
6.1111 | -102.0355
6.1148 | -102.0684
6.1185 | -102.107
6.1222 | -102.1491
6.1259 | -102.1873
6.1296 | -102.2218
6.1333 | -102.2544
6.137 | -102.2792
6.1407 | -102.2988
6.1444 | -102.302
6.1481 | -102.2942
6.1519 | -102.2857
6.1556 | -102.2816
6.1593 | -102.2773
6.163 | -102.2777
6.1667 | -102.28
6.1704 | -102.2874
6.1741 | -102.2987
6.1778 | -102.3161
6.1815 | -102.3399
6.1852 | -102.3716
6.1889 | -102.4039
6.1926 | -102.4331
6.1963 | -102.467
6.2 | -102.4955
6.2037 | -102.5112
6.2074 | -102.5208
6.2111 | -102.5211
6.2148 | -102.5106
6.2185 | -102.4907
6.2222 | -102.4635
6.2259 | -102.426
6.2296 | -102.3848
6.2333 | -102.355
6.237 | -102.3364
6.2407 | -102.3218
6.2444 | -102.3262
6.2481 | -102.3375
6.2519 | -102.3567
6.2556 | -102.384
6.2593 | -102.4194
6.263 | -102.4516
6.2667 | -102.4871
6.2704 | -102.5212
6.2741 | -102.5534
6.2778 | -102.5798
6.2815 | -102.5977
6.2852 | -102.6132
6.2889 | -102.6143
6.2926 | -102.6028
6.2963 | -102.5882
6.3 | -102.5622
6.3037 | -102.5233
6.3074 | -102.4782
6.3111 | -102.4425
6.3148 | -102.4108
6.3185 | -102.3875
6.3222 | -102.3672
6.3259 | -102.3507
6.3296 | -102.3469
6.3333 | -102.3416
6.337 | -102.3312
6.3407 | -102.3291
6.3444 | -102.3354
6.3481 | -102.3385
6.3519 | -102.3409
6.3556 | -102.3439
6.3593 | -102.3463
6.363 | -102.3548
6.3667 | -102.3596
6.3704 | -102.3671
6.3741 | -102.3756
6.3778 | -102.3778
6.3815 | -102.3861
6.3852 | -102.3957
6.3889 | -102.3984
6.3926 | -102.3972
6.3963 | -102.3931
6.4 | -102.3839
6.4037 | -102.3627
6.4074 | -102.3424
6.4111 | -102.3395
6.4148 | -102.3294
6.4185 | -102.324
6.4222 | -102.3219
6.4259 | -102.3235
6.4296 | -102.3376
6.4333 | -102.3602
6.437 | -102.3815
6.4407 | -102.4076
6.4444 | -102.4448
6.4481 | -102.4891
6.4519 | -102.5364
6.4556 | -102.5818
6.4593 | -102.6213
6.463 | -102.663
6.4667 | -102.7049
6.4704 | -102.7515
6.4741 | -102.7996
6.4778 | -102.8391
6.4815 | -102.8732
6.4852 | -102.9064
6.4889 | -102.9389
6.4926 | -102.9659
6.4963 | -102.9864
6.5 | -103.0002
6.5037 | -103.0118
6.5074 | -103.0089
6.5111 | -103.0051
6.5148 | -103.0023
6.5185 | -102.9922
6.5222 | -102.9772
6.5259 | -102.9514
6.5296 | -102.9236
6.5333 | -102.8862
6.537 | -102.8467
6.5407 | -102.7977
6.5444 | -102.7427
6.5481 | -102.6853
6.5519 | -102.6223
6.5556 | -102.5585
6.5593 | -102.5009
6.563 | -102.452
6.5667 | -102.4104
6.5704 | -102.3737
6.5741 | -102.3575
6.5778 | -102.3616
6.5815 | -102.3848
6.5852 | -102.4152
6.5889 | -102.4518
6.5926 | -102.5027
6.5963 | -102.5577
6.6 | -102.609
6.6037 | -102.652
6.6074 | -102.6884
6.6111 | -102.7177
6.6148 | -102.7413
6.6185 | -102.7535
6.6222 | -102.7535
6.6259 | -102.7507
6.6296 | -102.7381
6.6333 | -102.7225
6.637 | -102.7061
6.6407 | -102.6942
6.6444 | -102.6865
6.6481 | -102.6792
6.6519 | -102.6802
6.6556 | -102.6867
6.6593 | -102.6954
6.663 | -102.7096
6.6667 | -102.7334
6.6704 | -102.7636
6.6741 | -102.7882
6.6778 | -102.8139
6.6815 | -102.8377
6.6852 | -102.861
6.6889 | -102.877
6.6926 | -102.8863
6.6963 | -102.8962
6.7 | -102.9096
6.7037 | -102.9207
6.7074 | -102.9319
6.7111 | -102.9349
6.7148 | -102.9353
6.7185 | -102.9378
6.7222 | -102.9399
6.7259 | -102.9367
6.7296 | -102.9338
6.7333 | -102.9246
6.737 | -102.9049
6.7407 | -102.8831
6.7444 | -102.8574
6.7481 | -102.8316
6.7519 | -102.8046
6.7556 | -102.774
6.7593 | -102.7423
6.763 | -102.7093
6.7667 | -102.6746
6.7704 | -102.6494
6.7741 | -102.6302
6.7778 | -102.6172
6.7815 | -102.609
6.7852 | -102.6039
6.7889 | -102.6029
6.7926 | -102.6052
6.7963 | -102.6128
6.8 | -102.6258
6.8037 | -102.6511
6.8074 | -102.6822
6.8111 | -102.7135
6.8148 | -102.7507
6.8185 | -102.7916
6.8222 | -102.8342
6.8259 | -102.8799
6.8296 | -102.9163
6.8333 | -102.948
6.837 | -102.9781
6.8407 | -102.9981
6.8444 | -103.0119
6.8481 | -103.0142
6.8519 | -103.0103
6.8556 | -103.0119
6.8593 | -103.0085
6.863 | -102.9982
6.8667 | -102.9895
6.8704 | -102.9751
6.8741 | -102.9636
6.8778 | -102.9512
6.8815 | -102.9372
6.8852 | -102.9341
6.8889 | -102.9376
6.8926 | -102.9428
6.8963 | -102.9491
6.9 | -102.9583
6.9037 | -102.9672
6.9074 | -102.9767
6.9111 | -102.9967
6.9148 | -103.0126
6.9185 | -103.0316
6.9222 | -103.0521
6.9259 | -103.0738
6.9296 | -103.0981
6.9333 | -103.1234
6.937 | -103.1462
6.9407 | -103.1649
6.9444 | -103.1647
6.9481 | -103.1733
6.9519 | -103.1832
6.9556 | -103.182
6.9593 | -103.1739
6.963 | -103.1627
6.9667 | -103.1449
6.9704 | -103.1263
6.9741 | -103.109
6.9778 | -103.0896
6.9815 | -103.0657
6.9852 | -103.0417
6.9889 | -103.01
6.9926 | -102.9804
6.9963 | -102.9464
7 | -102.9117
7.0037 | -102.8817
7.0074 | -102.8535
7.0111 | -102.8267
7.0148 | -102.8019
7.0185 | -102.7867
7.0222 | -102.7801
7.0259 | -102.7795
7.0296 | -102.7862
7.0333 | -102.8037
7.037 | -102.8244
7.0407 | -102.8475
7.0444 | -102.8729
7.0481 | -102.9039
7.0519 | -102.9332
7.0556 | -102.9613
7.0593 | -102.9859
7.063 | -103.0089
7.0667 | -103.0263
7.0704 | -103.0416
7.0741 | -103.058
7.0778 | -103.0688
7.0815 | -103.0812
7.0852 | -103.0959
7.0889 | -103.1191
7.0926 | -103.1487
7.0963 | -103.1821
7.1 | -103.2184
7.1037 | -103.2718
7.1074 | -103.3279
7.1111 | -103.3844
7.1148 | -103.4425
7.1185 | -103.5033
7.1222 | -103.5545
7.1259 | -103.5981
7.1296 | -103.631
7.1333 | -103.6663
7.137 | -103.7
7.1407 | -103.7202
7.1444 | -103.7324
7.1481 | -103.7438
7.1519 | -103.7483
7.1556 | -103.7387
7.1593 | -103.7231
7.163 | -103.6984
7.1667 | -103.6661
7.1704 | -103.6251
7.1741 | -103.5789
7.1778 | -103.5293
7.1815 | -103.4745
7.1852 | -103.4171
7.1889 | -103.3612
7.1926 | -103.3124
7.1963 | -103.2692
7.2 | -103.2303
7.2037 | -103.1979
7.2074 | -103.1791
7.2111 | -103.1585
7.2148 | -103.1514
7.2185 | -103.152
7.2222 | -103.1577
7.2259 | -103.1674
7.2296 | -103.1861
7.2333 | -103.2057
7.237 | -103.229
7.2407 | -103.2466
7.2444 | -103.2659
7.2481 | -103.286
7.2519 | -103.3015
7.2556 | -103.3091
7.2593 | -103.3185
7.263 | -103.329
7.2667 | -103.3387
7.2704 | -103.3493
7.2741 | -103.3579
7.2778 | -103.37
7.2815 | -103.384
7.2852 | -103.3993
7.2889 | -103.4157
7.2926 | -103.4325
7.2963 | -103.4459
7.3 | -103.4574
7.3037 | -103.47
7.3074 | -103.483
7.3111 | -103.4868
7.3148 | -103.4741
7.3185 | -103.4647
7.3222 | -103.4446
7.3259 | -103.4168
7.3296 | -103.3875
7.3333 | -103.3585
7.337 | -103.328
7.3407 | -103.2964
7.3444 | -103.2691
7.3481 | -103.2461
7.3519 | -103.2235
7.3556 | -103.2145
7.3593 | -103.2155
7.363 | -103.2237
7.3667 | -103.239
7.3704 | -103.2656
7.3741 | -103.2973
7.3778 | -103.3355
7.3815 | -103.3812
7.3852 | -103.4308
7.3889 | -103.4806
7.3926 | -103.5314
7.3963 | -103.5825
7.4 | -103.6302
7.4037 | -103.6745
7.4074 | -103.708
7.4111 | -103.7274
7.4148 | -103.7369
7.4185 | -103.737
7.4222 | -103.7325
7.4259 | -103.7166
7.4296 | -103.7033
7.4333 | -103.6915
7.437 | -103.6747
7.4407 | -103.6582
7.4444 | -103.6405
7.4481 | -103.6252
7.4519 | -103.6055
7.4556 | -103.5814
7.4593 | -103.5602
7.463 | -103.5388
7.4667 | -103.5176
7.4704 | -103.5025
7.4741 | -103.487
7.4778 | -103.4932
7.4815 | -103.4933
7.4852 | -103.5043
7.4889 | -103.5222
7.4926 | -103.5399
7.4963 | -103.555
7.5 | -103.5677
7.5037 | -103.5816
7.5074 | -103.5947
7.5111 | -103.606
7.5148 | -103.6115
7.5185 | -103.6132
7.5222 | -103.615
7.5259 | -103.6092
7.5296 | -103.6035
7.5333 | -103.5989
7.537 | -103.5932
7.5407 | -103.579
7.5444 | -103.5694
7.5481 | -103.566
7.5519 | -103.5603
7.5556 | -103.5617
7.5593 | -103.5714
7.563 | -103.5843
7.5667 | -103.6016
7.5704 | -103.6176
7.5741 | -103.6298
7.5778 | -103.6437
7.5815 | -103.6597
7.5852 | -103.6777
7.5889 | -103.692
7.5926 | -103.6992
7.5963 | -103.6985
7.6 | -103.6911
7.6037 | -103.6794
7.6074 | -103.6746
7.6111 | -103.6687
7.6148 | -103.6612
7.6185 | -103.6492
7.6222 | -103.64
7.6259 | -103.6351
7.6296 | -103.6336
7.6333 | -103.6362
7.637 | -103.6393
7.6407 | -103.6522
7.6444 | -103.6676
7.6481 | -103.6788
7.6519 | -103.6967
7.6556 | -103.7243
7.6593 | -103.7496
7.663 | -103.7784
7.6667 | -103.8113
7.6704 | -103.8353
7.6741 | -103.8523
7.6778 | -103.8674
7.6815 | -103.8779
7.6852 | -103.877
7.6889 | -103.8757
7.6926 | -103.8703
7.6963 | -103.8553
7.7 | -103.8393
7.7037 | -103.8184
7.7074 | -103.8017
7.7111 | -103.7904
7.7148 | -103.7808
7.7185 | -103.772
7.7222 | -103.7605
7.7259 | -103.7505
7.7296 | -103.7462
7.7333 | -103.7469
7.737 | -103.7507
7.7407 | -103.7593
7.7444 | -103.772
7.7481 | -103.7813
7.7519 | -103.7919
7.7556 | -103.8005
7.7593 | -103.8056
7.763 | -103.8145
7.7667 | -103.8202
7.7704 | -103.8218
7.7741 | -103.8263
7.7778 | -103.8275
7.7815 | -103.8249
7.7852 | -103.8217
7.7889 | -103.8183
7.7926 | -103.8157
7.7963 | -103.8151
7.8 | -103.8195
7.8037 | -103.8221
7.8074 | -103.823
7.8111 | -103.8177
7.8148 | -103.8138
7.8185 | -103.8171
7.8222 | -103.8216
7.8259 | -103.8242
7.8296 | -103.8287
7.8333 | -103.8375
7.837 | -103.847
7.8407 | -103.8603
7.8444 | -103.8736
7.8481 | -103.8864
7.8519 | -103.9044
7.8556 | -103.9237
7.8593 | -103.9433
7.863 | -103.9694
7.8667 | -103.9967
7.8704 | -104.0191
7.8741 | -104.0442
7.8778 | -104.0645
7.8815 | -104.0781
7.8852 | -104.092
7.8889 | -104.1067
7.8926 | -104.1155
7.8963 | -104.1226
7.9 | -104.1238
7.9037 | -104.1079
7.9074 | -104.0843
7.9111 | -104.0505
7.9148 | -104.0089
7.9185 | -103.965
7.9222 | -103.9203
7.9259 | -103.8803
7.9296 | -103.832
7.9333 | -103.7856
7.937 | -103.7468
7.9407 | -103.7123
7.9444 | -103.6836
7.9481 | -103.6627
7.9519 | -103.6481
7.9556 | -103.6372
7.9593 | -103.6292
7.963 | -103.6268
7.9667 | -103.6237
7.9704 | -103.6233
7.9741 | -103.6258
7.9778 | -103.6269
7.9815 | -103.6243
7.9852 | -103.6245
7.9889 | -103.6234
7.9926 | -103.6194
7.9963 | -103.6195
8 | -103.6184
8.0037 | -103.6121
8.0074 | -103.6101
8.0111 | -103.609
8.0148 | -103.6181
8.0185 | -103.6214
8.0222 | -103.6311
8.0259 | -103.6439
8.0296 | -103.6622
8.0333 | -103.6821
8.037 | -103.7041
8.0407 | -103.7302
8.0444 | -103.757
8.0481 | -103.7877
8.0519 | -103.821
8.0556 | -103.8561
8.0593 | -103.8985
8.063 | -103.9421
8.0667 | -103.9864
8.0704 | -104.0334
8.0741 | -104.0806
8.0778 | -104.1279
8.0815 | -104.1705
8.0852 | -104.2127
8.0889 | -104.2545
8.0926 | -104.2937
8.0963 | -104.3259
8.1 | -104.3536
8.1037 | -104.3765
8.1074 | -104.3959
8.1111 | -104.4108
8.1148 | -104.4195
8.1185 | -104.4283
8.1222 | -104.4344
8.1259 | -104.4356
8.1296 | -104.4339
8.1333 | -104.4299
8.137 | -104.4267
8.1407 | -104.4188
8.1444 | -104.3982
8.1481 | -104.3674
8.1519 | -104.341
8.1556 | -104.3162
8.1593 | -104.2816
8.163 | -104.2465
8.1667 | -104.2129
8.1704 | -104.1714
8.1741 | -104.1261
8.1778 | -104.0793
8.1815 | -104.0365
8.1852 | -103.9969
8.1889 | -103.9583
8.1926 | -103.9281
8.1963 | -103.9038
8.2 | -103.8859
8.2037 | -103.8784
8.2074 | -103.8753
8.2111 | -103.8782
8.2148 | -103.8871
8.2185 | -103.8993
8.2222 | -103.9165
8.2259 | -103.935
8.2296 | -103.9541
8.2333 | -103.9745
8.237 | -103.9904
8.2407 | -103.9982
8.2444 | -104.0096
8.2481 | -104.0235
8.2519 | -104.0319
8.2556 | -104.0346
8.2593 | -104.0361
8.263 | -104.0333
8.2667 | -104.0313
8.2704 | -104.0268
8.2741 | -104.0117
8.2778 | -103.9944
8.2815 | -103.985
8.2852 | -103.9744
8.2889 | -103.9637
8.2926 | -103.9625
8.2963 | -103.9596
8.3 | -103.9666
8.3037 | -103.9778
8.3074 | -103.9914
8.3111 | -104.0103
8.3148 | -104.0288
8.3185 | -104.0525
8.3222 | -104.0783
8.3259 | -104.102
8.3296 | -104.1333
8.3333 | -104.1656
8.337 | -104.1919
8.3407 | -104.2188
8.3444 | -104.2502
8.3481 | -104.2793
8.3519 | -104.3076
8.3556 | -104.337
8.3593 | -104.3718
8.363 | -104.4019
8.3667 | -104.4292
8.3704 | -104.4575
8.3741 | -104.4785
8.3778 | -104.4927
8.3815 | -104.5052
8.3852 | -104.5191
8.3889 | -104.5292
8.3926 | -104.5276
8.3963 | -104.5184
8.4 | -104.51
8.4037 | -104.4945
8.4074 | -104.4801
8.4111 | -104.4643
8.4148 | -104.4476
8.4185 | -104.4353
8.4222 | -104.4174
8.4259 | -104.4053
8.4296 | -104.3975
8.4333 | -104.3921
8.437 | -104.3891
8.4407 | -104.3908
8.4444 | -104.3975
8.4481 | -104.4089
8.4519 | -104.4149
8.4556 | -104.4258
8.4593 | -104.4449
8.463 | -104.4687
8.4667 | -104.4855
8.4704 | -104.5027
8.4741 | -104.5205
8.4778 | -104.5363
8.4815 | -104.5451
8.4852 | -104.5586
8.4889 | -104.5684
8.4926 | -104.5697
8.4963 | -104.5763
8.5 | -104.5855
8.5037 | -104.5849
8.5074 | -104.5833
8.5111 | -104.5798
8.5148 | -104.5797
8.5185 | -104.5857
8.5222 | -104.5856
8.5259 | -104.5871
8.5296 | -104.59
8.5333 | -104.5998
8.537 | -104.603
8.5407 | -104.6057
8.5444 | -104.6116
8.5481 | -104.6071
8.5519 | -104.6076
8.5556 | -104.6087
8.5593 | -104.6037
8.563 | -104.5961
8.5667 | -104.5873
8.5704 | -104.5755
8.5741 | -104.5569
8.5778 | -104.5406
8.5815 | -104.522
8.5852 | -104.4993
8.5889 | -104.4739
8.5926 | -104.4458
8.5963 | -104.4159
8.6 | -104.3863
8.6037 | -104.3547
8.6074 | -104.3206
8.6111 | -104.2842
8.6148 | -104.2465
8.6185 | -104.2147
8.6222 | -104.1845
8.6259 | -104.1526
8.6296 | -104.1238
8.6333 | -104.0966
8.637 | -104.0774
8.6407 | -104.0643
8.6444 | -104.0507
8.6481 | -104.0447
8.6519 | -104.0449
8.6556 | -104.0477
8.6593 | -104.0521
8.663 | -104.0683
8.6667 | -104.0895
8.6704 | -104.114
8.6741 | -104.1444
8.6778 | -104.1761
8.6815 | -104.2097
8.6852 | -104.2472
8.6889 | -104.2848
8.6926 | -104.3242
8.6963 | -104.3568
8.7 | -104.3851
8.7037 | -104.415
8.7074 | -104.4386
8.7111 | -104.455
8.7148 | -104.464
8.7185 | -104.4663
8.7222 | -104.4631
8.7259 | -104.4548
8.7296 | -104.4403
8.7333 | -104.4245
8.737 | -104.4033
8.7407 | -104.3806
8.7444 | -104.3575
8.7481 | -104.3344
8.7519 | -104.3106
8.7556 | -104.2939
8.7593 | -104.2839
8.763 | -104.2777
8.7667 | -104.2744
8.7704 | -104.2747
8.7741 | -104.2812
8.7778 | -104.2967
8.7815 | -104.3207
8.7852 | -104.3495
8.7889 | -104.3845
8.7926 | -104.4191
8.7963 | -104.456
8.8 | -104.4941
8.8037 | -104.5304
8.8074 | -104.5644
8.8111 | -104.598
8.8148 | -104.6259
8.8185 | -104.6574
8.8222 | -104.6791
8.8259 | -104.6995
8.8296 | -104.7185
8.8333 | -104.7282
8.837 | -104.7345
8.8407 | -104.7402
8.8444 | -104.7449
8.8481 | -104.7472
8.8519 | -104.7487
8.8556 | -104.7511
8.8593 | -104.7525
8.863 | -104.7573
8.8667 | -104.7659
8.8704 | -104.7725
8.8741 | -104.7821
8.8778 | -104.7954
8.8815 | -104.8125
8.8852 | -104.8282
8.8889 | -104.8412
8.8926 | -104.8538
8.8963 | -104.8659
8.9 | -104.8765
8.9037 | -104.8876
8.9074 | -104.8929
8.9111 | -104.8926
8.9148 | -104.893
8.9185 | -104.8926
8.9222 | -104.8948
8.9259 | -104.894
8.9296 | -104.8897
8.9333 | -104.8858
8.937 | -104.8774
8.9407 | -104.8642
8.9444 | -104.8464
8.9481 | -104.8229
8.9519 | -104.8026
8.9556 | -104.7775
8.9593 | -104.7446
8.963 | -104.7112
8.9667 | -104.6751
8.9704 | -104.6372
8.9741 | -104.6004
8.9778 | -104.561
8.9815 | -104.527
8.9852 | -104.4952
8.9889 | -104.4646
8.9926 | -104.4355
8.9963 | -104.4173
9 | -104.4037
9.0037 | -104.3939
9.0074 | -104.3922
9.0111 | -104.3974
9.0148 | -104.409
9.0185 | -104.4275
9.0222 | -104.4506
9.0259 | -104.4751
9.0296 | -104.502
9.0333 | -104.5285
9.037 | -104.5562
9.0407 | -104.5855
9.0444 | -104.6133
9.0481 | -104.6326
9.0519 | -104.6473
9.0556 | -104.6608
9.0593 | -104.6685
9.063 | -104.6736
9.0667 | -104.6751
9.0704 | -104.6699
9.0741 | -104.6593
9.0778 | -104.6423
9.0815 | -104.6142
9.0852 | -104.5922
9.0889 | -104.5693
9.0926 | -104.5414
9.0963 | -104.5149
9.1 | -104.4937
9.1037 | -104.4715
9.1074 | -104.4494
9.1111 | -104.4355
9.1148 | -104.4268
9.1185 | -104.4194
9.1222 | -104.4168
9.1259 | -104.4194
9.1296 | -104.4226
9.1333 | -104.4318
9.137 | -104.4432
9.1407 | -104.4523
9.1444 | -104.4646
9.1481 | -104.4769
9.1519 | -104.4923
9.1556 | -104.5073
9.1593 | -104.5241
9.163 | -104.5413
9.1667 | -104.5575
9.1704 | -104.572
9.1741 | -104.5873
9.1778 | -104.6028
9.1815 | -104.621
9.1852 | -104.6334
9.1889 | -104.6399
9.1926 | -104.6502
9.1963 | -104.6564
9.2 | -104.6604
9.2037 | -104.6614
9.2074 | -104.662
9.2111 | -104.6614
9.2148 | -104.6589
9.2185 | -104.6605
9.2222 | -104.6622
9.2259 | -104.665
9.2296 | -104.6749
9.2333 | -104.6826
9.237 | -104.6851
9.2407 | -104.6899
9.2444 | -104.6958
9.2481 | -104.7005
9.2519 | -104.7048
9.2556 | -104.7111
9.2593 | -104.7173
9.263 | -104.7208
9.2667 | -104.724
9.2704 | -104.7199
9.2741 | -104.7195
9.2778 | -104.722
9.2815 | -104.7177
9.2852 | -104.7057
9.2889 | -104.6922
9.2926 | -104.6823
9.2963 | -104.6735
9.3 | -104.6614
9.3037 | -104.6436
9.3074 | -104.6293
9.3111 | -104.6152
9.3148 | -104.5991
9.3185 | -104.5832
9.3222 | -104.5683
9.3259 | -104.5547
9.3296 | -104.5363
9.3333 | -104.5209
9.337 | -104.5091
9.3407 | -104.4997
9.3444 | -104.4923
9.3481 | -104.4897
9.3519 | -104.4932
9.3556 | -104.4921
9.3593 | -104.4961
9.363 | -104.5103
9.3667 | -104.5293
9.3704 | -104.549
9.3741 | -104.5658
9.3778 | -104.5855
9.3815 | -104.6098
9.3852 | -104.6313
9.3889 | -104.6486
9.3926 | -104.6706
9.3963 | -104.6938
9.4 | -104.7159
9.4037 | -104.7357
9.4074 | -104.7593
9.4111 | -104.7777
9.4148 | -104.7928
9.4185 | -104.8126
9.4222 | -104.8371
9.4259 | -104.8599
9.4296 | -104.877
9.4333 | -104.8952
9.437 | -104.9178
9.4407 | -104.9455
9.4444 | -104.9687
9.4481 | -104.9928
9.4519 | -105.0175
9.4556 | -105.0429
9.4593 | -105.0691
9.463 | -105.093
9.4667 | -105.1128
9.4704 | -105.1241
9.4741 | -105.1279
9.4778 | -105.1334
9.4815 | -105.1334
9.4852 | -105.1262
9.4889 | -105.1125
9.4926 | -105.0986
9.4963 | -105.084
9.5 | -105.0632
9.5037 | -105.0452
9.5074 | -105.032
9.5111 | -105.0281
9.5148 | -105.0284
9.5185 | -105.0299
9.5222 | -105.0357
9.5259 | -105.0479
9.5296 | -105.0625
9.5333 | -105.0872
9.537 | -105.1173
9.5407 | -105.1512
9.5444 | -105.1828
9.5481 | -105.2194
9.5519 | -105.257
9.5556 | -105.2914
9.5593 | -105.3232
9.563 | -105.3471
9.5667 | -105.3709
9.5704 | -105.3888
9.5741 | -105.3934
9.5778 | -105.3936
9.5815 | -105.3863
9.5852 | -105.3719
9.5889 | -105.3509
9.5926 | -105.3216
9.5963 | -105.2843
9.6 | -105.2414
9.6037 | -105.1974
9.6074 | -105.1489
9.6111 | -105.1025
9.6148 | -105.0618
9.6185 | -105.0213
9.6222 | -104.9864
9.6259 | -104.9553
9.6296 | -104.9286
9.6333 | -104.9087
9.637 | -104.89
9.6407 | -104.8704
9.6444 | -104.849
9.6481 | -104.8322
9.6519 | -104.8194
9.6556 | -104.8097
9.6593 | -104.8015
9.663 | -104.7957
9.6667 | -104.7881
9.6704 | -104.7736
9.6741 | -104.755
9.6778 | -104.7368
9.6815 | -104.7171
9.6852 | -104.701
9.6889 | -104.6835
9.6926 | -104.67
9.6963 | -104.6598
9.7 | -104.6511
9.7037 | -104.6447
9.7074 | -104.6405
9.7111 | -104.639
9.7148 | -104.6452
9.7185 | -104.6532
9.7222 | -104.6628
9.7259 | -104.6776
9.7296 | -104.7009
9.7333 | -104.7277
9.737 | -104.7571
9.7407 | -104.7864
9.7444 | -104.8164
9.7481 | -104.8475
9.7519 | -104.8745
9.7556 | -104.9015
9.7593 | -104.9233
9.763 | -104.9438
9.7667 | -104.957
9.7704 | -104.965
9.7741 | -104.9727
9.7778 | -104.9776
9.7815 | -104.9783
9.7852 | -104.9786
9.7889 | -104.9814
9.7926 | -104.9754
9.7963 | -104.9613
9.8 | -104.9397
9.8037 | -104.9162
9.8074 | -104.89
9.8111 | -104.8691
9.8148 | -104.8522
9.8185 | -104.8294
9.8222 | -104.804
9.8259 | -104.7809
9.8296 | -104.7563
9.8333 | -104.7312
9.837 | -104.7074
9.8407 | -104.6855
9.8444 | -104.6655
9.8481 | -104.6517
9.8519 | -104.6435
9.8556 | -104.6409
9.8593 | -104.6414
9.863 | -104.6443
9.8667 | -104.6492
9.8704 | -104.6568
9.8741 | -104.6646
9.8778 | -104.6753
9.8815 | -104.6878
9.8852 | -104.7078
9.8889 | -104.726
9.8926 | -104.7438
9.8963 | -104.7618
9.9 | -104.7797
9.9037 | -104.7991
9.9074 | -104.8126
9.9111 | -104.8276
9.9148 | -104.8427
9.9185 | -104.8542
9.9222 | -104.8653
9.9259 | -104.8781
9.9296 | -104.8892
9.9333 | -104.8961
9.937 | -104.9086
9.9407 | -104.9235
9.9444 | -104.9412
9.9481 | -104.9647
9.9519 | -104.9918
9.9556 | -105.0167
9.9593 | -105.0418
9.963 | -105.0745
9.9667 | -105.1095
9.9704 | -105.1439
9.9741 | -105.1794
9.9778 | -105.2131
9.9815 | -105.2412
9.9852 | -105.2719
9.9889 | -105.3044
9.9926 | -105.331
9.9963 | -105.3487
10 | -105.3605
10.0037 | -105.3689
10.0074 | -105.3706
10.0111 | -105.3713
10.0148 | -105.3611
10.0185 | -105.3443
10.0222 | -105.3281
10.0259 | -105.3054
10.0296 | -105.2751
10.0333 | -105.249
10.037 | -105.2251
10.0407 | -105.1993
10.0444 | -105.1731
10.0481 | -105.1525
10.0519 | -105.1291
10.0556 | -105.1081
10.0593 | -105.0982
10.063 | -105.0932
10.0667 | -105.0902
10.0704 | -105.0906
10.0741 | -105.0978
10.0778 | -105.1082
10.0815 | -105.1167
10.0852 | -105.1251
10.0889 | -105.133
10.0926 | -105.136
10.0963 | -105.1377
10.1 | -105.1366
10.1037 | -105.1371
10.1074 | -105.1314
10.1111 | -105.1234
10.1148 | -105.1144
10.1185 | -105.1008
10.1222 | -105.0846
10.1259 | -105.0691
10.1296 | -105.0515
10.1333 | -105.0288
10.137 | -105.0069
10.1407 | -104.9858
10.1444 | -104.966
10.1481 | -104.9427
10.1519 | -104.9255
10.1556 | -104.9063
10.1593 | -104.8873
10.163 | -104.872
10.1667 | -104.8596
10.1704 | -104.8462
10.1741 | -104.8373
10.1778 | -104.8294
10.1815 | -104.8194
10.1852 | -104.8089
10.1889 | -104.8051
10.1926 | -104.803
10.1963 | -104.8034
10.2 | -104.8136
10.2037 | -104.8242
10.2074 | -104.8283
10.2111 | -104.8362
10.2148 | -104.8463
10.2185 | -104.8611
10.2222 | -104.8758
10.2259 | -104.8913
10.2296 | -104.9101
10.2333 | -104.9305
10.237 | -104.9533
10.2407 | -104.9755
10.2444 | -104.9997
10.2481 | -105.0254
10.2519 | -105.0478
10.2556 | -105.0691
10.2593 | -105.0901
10.263 | -105.11
10.2667 | -105.1289
10.2704 | -105.1494
10.2741 | -105.1715
10.2778 | -105.192
10.2815 | -105.213
10.2852 | -105.2321
10.2889 | -105.248
10.2926 | -105.2608
10.2963 | -105.2752
10.3 | -105.2836
10.3037 | -105.2936
10.3074 | -105.302
10.3111 | -105.3072
10.3148 | -105.3091
10.3185 | -105.3073
10.3222 | -105.3041
10.3259 | -105.2992
10.3296 | -105.2918
10.3333 | -105.2835
10.337 | -105.2719
10.3407 | -105.2555
10.3444 | -105.2383
10.3481 | -105.2186
10.3519 | -105.1944
10.3556 | -105.1712
10.3593 | -105.1524
10.363 | -105.1292
10.3667 | -105.1089
10.3704 | -105.0923
10.3741 | -105.0788
10.3778 | -105.0686
10.3815 | -105.0599
10.3852 | -105.053
10.3889 | -105.0481
10.3926 | -105.0442
10.3963 | -105.04
10.4 | -105.0413
10.4037 | -105.0487
10.4074 | -105.0618
10.4111 | -105.077
10.4148 | -105.0916
10.4185 | -105.1085
10.4222 | -105.127
10.4259 | -105.1448
10.4296 | -105.1637
10.4333 | -105.1834
10.437 | -105.2055
10.4407 | -105.2305
10.4444 | -105.2543
10.4481 | -105.2757
10.4519 | -105.2969
10.4556 | -105.3128
10.4593 | -105.3312
10.463 | -105.3497
10.4667 | -105.3659
10.4704 | -105.3807
10.4741 | -105.3944
10.4778 | -105.4075
10.4815 | -105.4155
10.4852 | -105.4241
10.4889 | -105.4339
10.4926 | -105.4413
10.4963 | -105.4491
10.5 | -105.4561
10.5037 | -105.4646
10.5074 | -105.4717
10.5111 | -105.4741
10.5148 | -105.4778
10.5185 | -105.4785
10.5222 | -105.4746
10.5259 | -105.4731
10.5296 | -105.4703
10.5333 | -105.4662
10.537 | -105.4621
10.5407 | -105.4532
10.5444 | -105.4464
10.5481 | -105.4385
10.5519 | -105.4187
10.5556 | -105.3971
10.5593 | -105.3765
10.563 | -105.3548
10.5667 | -105.3325
10.5704 | -105.3125
10.5741 | -105.2937
10.5778 | -105.2761
10.5815 | -105.2616
10.5852 | -105.2521
10.5889 | -105.2446
10.5926 | -105.2388
10.5963 | -105.2306
10.6 | -105.2269
10.6037 | -105.2292
10.6074 | -105.2357
10.6111 | -105.243
10.6148 | -105.2526
10.6185 | -105.2679
10.6222 | -105.2849
10.6259 | -105.3034
10.6296 | -105.3224
10.6333 | -105.3432
10.637 | -105.3631
10.6407 | -105.3817
10.6444 | -105.4008
10.6481 | -105.4162
10.6519 | -105.4326
10.6556 | -105.4472
10.6593 | -105.4595
10.663 | -105.4743
10.6667 | -105.4872
10.6704 | -105.4979
10.6741 | -105.5103
10.6778 | -105.5199
10.6815 | -105.5279
10.6852 | -105.5403
10.6889 | -105.5523
10.6926 | -105.5615
10.6963 | -105.5704
10.7 | -105.578
10.7037 | -105.5851
10.7074 | -105.5994
10.7111 | -105.6114
10.7148 | -105.6211
10.7185 | -105.6301
10.7222 | -105.637
10.7259 | -105.6429
10.7296 | -105.6435
10.7333 | -105.6419
10.737 | -105.6382
10.7407 | -105.6336
10.7444 | -105.6277
10.7481 | -105.6227
10.7519 | -105.6221
10.7556 | -105.6145
10.7593 | -105.6032
10.763 | -105.592
10.7667 | -105.5786
10.7704 | -105.5658
10.7741 | -105.5545
10.7778 | -105.5419
10.7815 | -105.5258
10.7852 | -105.5149
10.7889 | -105.5067
10.7926 | -105.4933
10.7963 | -105.4775
10.8 | -105.4596
10.8037 | -105.4434
10.8074 | -105.4259
10.8111 | -105.4126
10.8148 | -105.4014
10.8185 | -105.3887
10.8222 | -105.3782
10.8259 | -105.3679
10.8296 | -105.3572
10.8333 | -105.3479
10.837 | -105.3412
10.8407 | -105.3333
10.8444 | -105.3261
10.8481 | -105.3192
10.8519 | -105.3141
10.8556 | -105.3115
10.8593 | -105.3059
10.863 | -105.3024
10.8667 | -105.2963
10.8704 | -105.2905
10.8741 | -105.2814
10.8778 | -105.273
10.8815 | -105.2653
10.8852 | -105.2542
10.8889 | -105.2452
10.8926 | -105.2325
10.8963 | -105.2199
10.9 | -105.2075
10.9037 | -105.1956
10.9074 | -105.1876
10.9111 | -105.1793
10.9148 | -105.1764
10.9185 | -105.1748
10.9222 | -105.1724
10.9259 | -105.176
10.9296 | -105.1807
10.9333 | -105.1855
10.937 | -105.1954
10.9407 | -105.2086
10.9444 | -105.219
10.9481 | -105.2314
10.9519 | -105.2497
10.9556 | -105.2725
10.9593 | -105.295
10.963 | -105.3158
10.9667 | -105.3318
10.9704 | -105.3461
10.9741 | -105.3623
10.9778 | -105.3799
10.9815 | -105.3961
10.9852 | -105.4154
10.9889 | -105.4318
10.9926 | -105.4468
10.9963 | -105.4643
11 | -105.4797
11.0037 | -105.4897
11.0074 | -105.4994
11.0111 | -105.5049
11.0148 | -105.5053
11.0185 | -105.5046
11.0222 | -105.4987
11.0259 | -105.4922
11.0296 | -105.4811
11.0333 | -105.4701
11.037 | -105.4563
11.0407 | -105.4434
11.0444 | -105.4265
11.0481 | -105.4036
11.0519 | -105.3852
11.0556 | -105.3652
11.0593 | -105.3448
11.063 | -105.3246
11.0667 | -105.303
11.0704 | -105.2865
11.0741 | -105.2703
11.0778 | -105.255
11.0815 | -105.2451
11.0852 | -105.2346
11.0889 | -105.2263
11.0926 | -105.221
11.0963 | -105.2186
11.1 | -105.2186
11.1037 | -105.2227
11.1074 | -105.2281
11.1111 | -105.2336
11.1148 | -105.241
11.1185 | -105.251
11.1222 | -105.2603
11.1259 | -105.2762
11.1296 | -105.2927
11.1333 | -105.3067
11.137 | -105.3202
11.1407 | -105.3335
11.1444 | -105.3446
11.1481 | -105.3595
11.1519 | -105.3765
11.1556 | -105.3889
11.1593 | -105.4049
11.163 | -105.4222
11.1667 | -105.4368
11.1704 | -105.4533
11.1741 | -105.4701
11.1778 | -105.4861
11.1815 | -105.4978
11.1852 | -105.5047
11.1889 | -105.5164
11.1926 | -105.5302
11.1963 | -105.539
11.2 | -105.5431
11.2037 | -105.5515
11.2074 | -105.5607
11.2111 | -105.5635
11.2148 | -105.5637
11.2185 | -105.5603
11.2222 | -105.5594
11.2259 | -105.5566
11.2296 | -105.5552
11.2333 | -105.552
11.237 | -105.55
11.2407 | -105.5507
11.2444 | -105.552
11.2481 | -105.5551
11.2519 | -105.556
11.2556 | -105.5588
11.2593 | -105.5664
11.263 | -105.5743
11.2667 | -105.5791
11.2704 | -105.5838
11.2741 | -105.5935
11.2778 | -105.6054
11.2815 | -105.6164
11.2852 | -105.6283
11.2889 | -105.6353
11.2926 | -105.6403
11.2963 | -105.6433
11.3 | -105.6469
11.3037 | -105.6503
11.3074 | -105.6547
11.3111 | -105.6541
11.3148 | -105.6483
11.3185 | -105.6458
11.3222 | -105.6392
11.3259 | -105.6353
11.3296 | -105.6321
11.3333 | -105.6249
11.337 | -105.6191
11.3407 | -105.6107
11.3444 | -105.602
11.3481 | -105.5974
11.3519 | -105.59
11.3556 | -105.5834
11.3593 | -105.5799
11.363 | -105.5782
11.3667 | -105.5757
11.3704 | -105.5787
11.3741 | -105.5852
11.3778 | -105.5935
11.3815 | -105.6068
11.3852 | -105.6225
11.3889 | -105.6389
11.3926 | -105.6578
11.3963 | -105.6748
11.4 | -105.6936
11.4037 | -105.7158
11.4074 | -105.7391
11.4111 | -105.761
11.4148 | -105.7782
11.4185 | -105.7955
11.4222 | -105.8121
11.4259 | -105.8247
11.4296 | -105.8369
11.4333 | -105.848
11.437 | -105.8581
11.4407 | -105.8611
11.4444 | -105.8591
11.4481 | -105.8587
11.4519 | -105.8562
11.4556 | -105.8528
11.4593 | -105.848
11.463 | -105.839
11.4667 | -105.8272
11.4704 | -105.8117
11.4741 | -105.7926
11.4778 | -105.7722
11.4815 | -105.7546
11.4852 | -105.735
11.4889 | -105.7146
11.4926 | -105.6902
11.4963 | -105.6683
11.5 | -105.6517
11.5037 | -105.6305
11.5074 | -105.6136
11.5111 | -105.6005
11.5148 | -105.5865
11.5185 | -105.5721
11.5222 | -105.5559
11.5259 | -105.5476
11.5296 | -105.5397
11.5333 | -105.5332
11.537 | -105.527
11.5407 | -105.5232
11.5444 | -105.5184
11.5481 | -105.5127
11.5519 | -105.5084
11.5556 | -105.5065
11.5593 | -105.5062
11.563 | -105.5091
11.5667 | -105.5109
11.5704 | -105.5084
11.5741 | -105.507
11.5778 | -105.5086
11.5815 | -105.5123
11.5852 | -105.5136
11.5889 | -105.511
11.5926 | -105.5066
11.5963 | -105.5002
11.6 | -105.4936
11.6037 | -105.4853
11.6074 | -105.4786
11.6111 | -105.4697
11.6148 | -105.4583
11.6185 | -105.4463
11.6222 | -105.4334
11.6259 | -105.4227
11.6296 | -105.4162
11.6333 | -105.409
11.637 | -105.401
11.6407 | -105.3947
11.6444 | -105.3918
11.6481 | -105.3918
11.6519 | -105.3916
11.6556 | -105.3944
11.6593 | -105.3966
11.663 | -105.3992
11.6667 | -105.4052
11.6704 | -105.4171
11.6741 | -105.4293
11.6778 | -105.4425
11.6815 | -105.4543
11.6852 | -105.4674
11.6889 | -105.4807
11.6926 | -105.4938
11.6963 | -105.507
11.7 | -105.5186
11.7037 | -105.5313
11.7074 | -105.5453
11.7111 | -105.5623
11.7148 | -105.5786
11.7185 | -105.5916
11.7222 | -105.6048
11.7259 | -105.6184
11.7296 | -105.629
11.7333 | -105.6403
11.737 | -105.6502
11.7407 | -105.6585
11.7444 | -105.6663
11.7481 | -105.6744
11.7519 | -105.6785
11.7556 | -105.6876
11.7593 | -105.6963
11.763 | -105.7042
11.7667 | -105.7112
11.7704 | -105.7174
11.7741 | -105.7219
11.7778 | -105.7222
11.7815 | -105.7238
11.7852 | -105.7254
11.7889 | -105.7271
11.7926 | -105.7305
11.7963 | -105.7318
11.8 | -105.7324
11.8037 | -105.7309
11.8074 | -105.7319
11.8111 | -105.7333
11.8148 | -105.73
11.8185 | -105.7264
11.8222 | -105.7272
11.8259 | -105.7283
11.8296 | -105.7305
11.8333 | -105.7346
11.837 | -105.7373
11.8407 | -105.7418
11.8444 | -105.7473
11.8481 | -105.7574
11.8519 | -105.7696
11.8556 | -105.7821
11.8593 | -105.8007
11.863 | -105.8181
11.8667 | -105.8323
11.8704 | -105.8461
11.8741 | -105.8611
11.8778 | -105.8754
11.8815 | -105.8896
11.8852 | -105.901
11.8889 | -105.9119
11.8926 | -105.9189
11.8963 | -105.9237
11.9 | -105.9288
11.9037 | -105.9329
11.9074 | -105.9386
11.9111 | -105.9416
11.9148 | -105.9384
11.9185 | -105.9311
11.9222 | -105.9215
11.9259 | -105.9105
11.9296 | -105.8956
11.9333 | -105.8792
11.937 | -105.8616
11.9407 | -105.8451
11.9444 | -105.8289
11.9481 | -105.8087
11.9519 | -105.7881
11.9556 | -105.7651
11.9593 | -105.7381
11.963 | -105.7057
11.9667 | -105.6729
11.9704 | -105.6419
11.9741 | -105.6128
11.9778 | -105.5891
11.9815 | -105.5669
11.9852 | -105.546
11.9889 | -105.5273
11.9926 | -105.5078
11.9963 | -105.4916
12 | -105.4772
12.0037 | -105.4676
12.0074 | -105.4601
12.0111 | -105.4567
12.0148 | -105.454
12.0185 | -105.4544
12.0222 | -105.4588
12.0259 | -105.467
12.0296 | -105.478
12.0333 | -105.4886
12.037 | -105.5004
12.0407 | -105.5153
12.0444 | -105.5309
12.0481 | -105.5456
12.0519 | -105.5614
12.0556 | -105.5765
12.0593 | -105.5963
12.063 | -105.617
12.0667 | -105.6369
12.0704 | -105.6531
12.0741 | -105.6698
12.0778 | -105.6859
12.0815 | -105.6996
12.0852 | -105.7124
12.0889 | -105.7275
12.0926 | -105.7463
12.0963 | -105.7647
12.1 | -105.7793
12.1037 | -105.7938
12.1074 | -105.8037
12.1111 | -105.8134
12.1148 | -105.8271
12.1185 | -105.8391
12.1222 | -105.853
12.1259 | -105.8638
12.1296 | -105.8738
12.1333 | -105.8868
12.137 | -105.9038
12.1407 | -105.9176
12.1444 | -105.9302
12.1481 | -105.9391
12.1519 | -105.9435
12.1556 | -105.9487
12.1593 | -105.9517
12.163 | -105.9575
12.1667 | -105.9623
12.1704 | -105.9623
12.1741 | -105.9669
12.1778 | -105.9684
12.1815 | -105.9647
12.1852 | -105.9577
12.1889 | -105.9525
12.1926 | -105.9486
12.1963 | -105.9412
12.2 | -105.9355
12.2037 | -105.93
12.2074 | -105.9232
12.2111 | -105.9182
12.2148 | -105.919
12.2185 | -105.9214
12.2222 | -105.9206
12.2259 | -105.9187
12.2296 | -105.9182
12.2333 | -105.9147
12.237 | -105.913
12.2407 | -105.9148
12.2444 | -105.9155
12.2481 | -105.9177
12.2519 | -105.9184
12.2556 | -105.919
12.2593 | -105.9207
12.263 | -105.9238
12.2667 | -105.9275
12.2704 | -105.9321
12.2741 | -105.9347
12.2778 | -105.9346
12.2815 | -105.9372
12.2852 | -105.9434
12.2889 | -105.944
12.2926 | -105.9405
12.2963 | -105.9387
12.3 | -105.9397
12.3037 | -105.9429
12.3074 | -105.9457
12.3111 | -105.9433
12.3148 | -105.9382
12.3185 | -105.9322
12.3222 | -105.9233
12.3259 | -105.9174
12.3296 | -105.9108
12.3333 | -105.8982
12.337 | -105.8835
12.3407 | -105.8675
12.3444 | -105.8517
12.3481 | -105.8386
12.3519 | -105.8197
12.3556 | -105.8003
12.3593 | -105.7832
12.363 | -105.7662
12.3667 | -105.7487
12.3704 | -105.7343
12.3741 | -105.722
12.3778 | -105.7079
12.3815 | -105.6942
12.3852 | -105.6841
12.3889 | -105.6734
12.3926 | -105.6633
12.3963 | -105.6556
12.4 | -105.6522
12.4037 | -105.6542
12.4074 | -105.6574
12.4111 | -105.6621
12.4148 | -105.6651
12.4185 | -105.6695
12.4222 | -105.6775
12.4259 | -105.6822
12.4296 | -105.6895
12.4333 | -105.6985
12.437 | -105.7083
12.4407 | -105.7176
12.4444 | -105.7252
12.4481 | -105.7334
12.4519 | -105.7444
12.4556 | -105.7571
12.4593 | -105.7685
12.463 | -105.7753
12.4667 | -105.7847
12.4704 | -105.7951
12.4741 | -105.8053
12.4778 | -105.8128
12.4815 | -105.8135
12.4852 | -105.8168
12.4889 | -105.8225
12.4926 | -105.8293
12.4963 | -105.8338
12.5 | -105.8391
12.5037 | -105.8447
12.5074 | -105.8483
12.5111 | -105.8517
12.5148 | -105.8557
12.5185 | -105.8611
12.5222 | -105.868
12.5259 | -105.874
12.5296 | -105.8774
12.5333 | -105.8843
12.537 | -105.8933
12.5407 | -105.902
12.5444 | -105.9102
12.5481 | -105.919
12.5519 | -105.9277
12.5556 | -105.9374
12.5593 | -105.9511
12.563 | -105.964
12.5667 | -105.9719
12.5704 | -105.9819
12.5741 | -105.9929
12.5778 | -106.0011
12.5815 | -106.0061
12.5852 | -106.0098
12.5889 | -106.0129
12.5926 | -106.0149
12.5963 | -106.0191
12.6 | -106.0201
12.6037 | -106.0173
12.6074 | -106.0116
12.6111 | -106.0096
12.6148 | -106.0045
12.6185 | -105.9951
12.6222 | -105.9851
12.6259 | -105.9751
12.6296 | -105.9672
12.6333 | -105.9572
12.637 | -105.9469
12.6407 | -105.9382
12.6444 | -105.9312
12.6481 | -105.9236
12.6519 | -105.916
12.6556 | -105.9073
12.6593 | -105.9004
12.663 | -105.898
12.6667 | -105.8917
12.6704 | -105.8802
12.6741 | -105.8701
12.6778 | -105.8646
12.6815 | -105.8597
12.6852 | -105.858
12.6889 | -105.8586
12.6926 | -105.8577
12.6963 | -105.8592
12.7 | -105.8593
12.7037 | -105.8536
12.7074 | -105.8511
12.7111 | -105.8508
12.7148 | -105.851
12.7185 | -105.8506
12.7222 | -105.8489
12.7259 | -105.8452
12.7296 | -105.8399
12.7333 | -105.8303
12.737 | -105.8177
12.7407 | -105.8056
12.7444 | -105.7944
12.7481 | -105.7843
12.7519 | -105.7707
12.7556 | -105.7556
12.7593 | -105.7425
12.763 | -105.7286
12.7667 | -105.716
12.7704 | -105.7044
12.7741 | -105.6946
12.7778 | -105.6862
12.7815 | -105.6828
12.7852 | -105.6819
12.7889 | -105.682
12.7926 | -105.6818
12.7963 | -105.6858
12.8 | -105.6908
12.8037 | -105.6993
12.8074 | -105.7109
12.8111 | -105.7236
12.8148 | -105.7368
12.8185 | -105.7524
12.8222 | -105.7676
12.8259 | -105.7803
12.8296 | -105.7922
12.8333 | -105.8034
12.837 | -105.8157
12.8407 | -105.8274
12.8444 | -105.8372
12.8481 | -105.8481
12.8519 | -105.8591
12.8556 | -105.8715
12.8593 | -105.8829
12.863 | -105.8924
12.8667 | -105.8974
12.8704 | -105.9002
12.8741 | -105.9036
12.8778 | -105.9095
12.8815 | -105.9148
12.8852 | -105.9244
12.8889 | -105.939
12.8926 | -105.9502
12.8963 | -105.9585
12.9 | -105.9684
12.9037 | -105.979
12.9074 | -105.9907
12.9111 | -106.005
12.9148 | -106.0164
12.9185 | -106.029
12.9222 | -106.0447
12.9259 | -106.0588
12.9296 | -106.0733
12.9333 | -106.0878
12.937 | -106.1014
12.9407 | -106.1146
12.9444 | -106.127
12.9481 | -106.1432
12.9519 | -106.159
12.9556 | -106.1762
12.9593 | -106.189
12.963 | -106.199
12.9667 | -106.2096
12.9704 | -106.2217
12.9741 | -106.2354
12.9778 | -106.2488
12.9815 | -106.2594
12.9852 | -106.2718
12.9889 | -106.2824
12.9926 | -106.2913
12.9963 | -106.2979
13 | -106.3054
13.0037 | -106.3086
13.0074 | -106.3115
13.0111 | -106.3132
13.0148 | -106.3121
13.0185 | -106.3115
13.0222 | -106.3112
13.0259 | -106.3092
13.0296 | -106.3054
13.0333 | -106.3022
13.037 | -106.2946
13.0407 | -106.2868
13.0444 | -106.2807
13.0481 | -106.2742
13.0519 | -106.2664
13.0556 | -106.2621
13.0593 | -106.2598
13.063 | -106.258
13.0667 | -106.2577
13.0704 | -106.255
13.0741 | -106.2509
13.0778 | -106.2465
13.0815 | -106.2452
13.0852 | -106.2469
13.0889 | -106.2495
13.0926 | -106.2543
13.0963 | -106.2605
13.1 | -106.2638
13.1037 | -106.2677
13.1074 | -106.2731
13.1111 | -106.2769
13.1148 | -106.2804
13.1185 | -106.2835
13.1222 | -106.2854
13.1259 | -106.2912
13.1296 | -106.2937
13.1333 | -106.2965
13.137 | -106.2955
13.1407 | -106.2919
13.1444 | -106.2897
13.1481 | -106.2891
13.1519 | -106.2874
13.1556 | -106.2845
13.1593 | -106.2771
13.163 | -106.2689
13.1667 | -106.2608
13.1704 | -106.2528
13.1741 | -106.243
13.1778 | -106.2334
13.1815 | -106.2234
13.1852 | -106.2136
13.1889 | -106.2065
13.1926 | -106.2003
13.1963 | -106.1949
13.2 | -106.192
13.2037 | -106.1903
13.2074 | -106.1847
13.2111 | -106.181
13.2148 | -106.178
13.2185 | -106.1763
13.2222 | -106.1739
13.2259 | -106.1732
13.2296 | -106.173
13.2333 | -106.1741
13.237 | -106.175
13.2407 | -106.1756
13.2444 | -106.1765
13.2481 | -106.1728
13.2519 | -106.1677
13.2556 | -106.1665
13.2593 | -106.165
13.263 | -106.1622
13.2667 | -106.1594
13.2704 | -106.157
13.2741 | -106.1554
13.2778 | -106.155
13.2815 | -106.1553
13.2852 | -106.1577
13.2889 | -106.1615
13.2926 | -106.1656
13.2963 | -106.1683
13.3 | -106.1721
13.3037 | -106.1767
13.3074 | -106.1813
13.3111 | -106.1869
13.3148 | -106.1888
13.3185 | -106.1938
13.3222 | -106.2018
13.3259 | -106.2105
13.3296 | -106.2158
13.3333 | -106.2214
13.337 | -106.229
13.3407 | -106.2333
13.3444 | -106.236
13.3481 | -106.2431
13.3519 | -106.25
13.3556 | -106.2515
13.3593 | -106.2574
13.363 | -106.2631
13.3667 | -106.2659
13.3704 | -106.269
13.3741 | -106.2759
13.3778 | -106.2817
13.3815 | -106.2838
13.3852 | -106.2838
13.3889 | -106.2834
13.3926 | -106.282
13.3963 | -106.2832
13.4 | -106.2856
13.4037 | -106.2865
13.4074 | -106.2863
13.4111 | -106.2868
13.4148 | -106.2875
13.4185 | -106.2864
13.4222 | -106.2861
13.4259 | -106.2864
13.4296 | -106.2871
13.4333 | -106.2897
13.437 | -106.2924
13.4407 | -106.2964
13.4444 | -106.3017
13.4481 | -106.3024
13.4519 | -106.3055
13.4556 | -106.3077
13.4593 | -106.3081
13.463 | -106.3077
13.4667 | -106.3078
13.4704 | -106.3065
13.4741 | -106.3033
13.4778 | -106.3016
13.4815 | -106.2997
13.4852 | -106.2992
13.4889 | -106.2991
13.4926 | -106.2968
13.4963 | -106.2955
13.5 | -106.2923
13.5037 | -106.2889
13.5074 | -106.2862
13.5111 | -106.2819
13.5148 | -106.2754
13.5185 | -106.2657
13.5222 | -106.2615
13.5259 | -106.2539
13.5296 | -106.2457
13.5333 | -106.2419
13.537 | -106.2352
13.5407 | -106.2286
13.5444 | -106.2229
13.5481 | -106.2205
13.5519 | -106.2175
13.5556 | -106.2097
13.5593 | -106.2043
13.563 | -106.2047
13.5667 | -106.2056
13.5704 | -106.2023
13.5741 | -106.2026
13.5778 | -106.2058
13.5815 | -106.2073
13.5852 | -106.2066
13.5889 | -106.2058
13.5926 | -106.2067
13.5963 | -106.2088
13.6 | -106.2095
13.6037 | -106.212
13.6074 | -106.2148
13.6111 | -106.2187
13.6148 | -106.2232
13.6185 | -106.2268
13.6222 | -106.2279
13.6259 | -106.2274
13.6296 | -106.2281
13.6333 | -106.229
13.637 | -106.2328
13.6407 | -106.2356
13.6444 | -106.2382
13.6481 | -106.2396
13.6519 | -106.2425
13.6556 | -106.2441
13.6593 | -106.2447
13.663 | -106.2464
13.6667 | -106.2459
13.6704 | -106.244
13.6741 | -106.2416
13.6778 | -106.2418
13.6815 | -106.2413
13.6852 | -106.2373
13.6889 | -106.2339
13.6926 | -106.2292
13.6963 | -106.2267
13.7 | -106.2232
13.7037 | -106.2242
13.7074 | -106.2254
13.7111 | -106.224
13.7148 | -106.2241
13.7185 | -106.2242
13.7222 | -106.2266
13.7259 | -106.2291
13.7296 | -106.2323
13.7333 | -106.2358
13.737 | -106.238
13.7407 | -106.239
13.7444 | -106.2425
13.7481 | -106.2474
13.7519 | -106.2553
13.7556 | -106.2637
13.7593 | -106.2712
13.763 | -106.2765
13.7667 | -106.2864
13.7704 | -106.2975
13.7741 | -106.3084
13.7778 | -106.3234
13.7815 | -106.3376
13.7852 | -106.3465
13.7889 | -106.3545
13.7926 | -106.3641
13.7963 | -106.3746
13.8 | -106.3836
13.8037 | -106.3924
13.8074 | -106.4001
13.8111 | -106.4089
13.8148 | -106.4169
13.8185 | -106.4248
13.8222 | -106.436
13.8259 | -106.4419
13.8296 | -106.4469
13.8333 | -106.4516
13.837 | -106.4528
13.8407 | -106.4543
13.8444 | -106.4567
13.8481 | -106.4577
13.8519 | -106.4603
13.8556 | -106.4631
13.8593 | -106.4635
13.863 | -106.468
13.8667 | -106.475
13.8704 | -106.4812
13.8741 | -106.4866
13.8778 | -106.4934
13.8815 | -106.5034
13.8852 | -106.5158
13.8889 | -106.5364
13.8926 | -106.5505
13.8963 | -106.5656
13.9 | -106.5797
13.9037 | -106.5886
13.9074 | -106.6015
13.9111 | -106.6145
13.9148 | -106.6287
13.9185 | -106.6419
13.9222 | -106.6528
13.9259 | -106.6629
13.9296 | -106.6702
13.9333 | -106.6744
13.937 | -106.6798
13.9407 | -106.6854
13.9444 | -106.6847
13.9481 | -106.6837
13.9519 | -106.6777
13.9556 | -106.6688
13.9593 | -106.6477
13.963 | -106.6188
13.9667 | -106.5909
13.9704 | -106.5587
13.9741 | -106.5252
13.9778 | -106.4933
13.9815 | -106.4618
13.9852 | -106.4293
13.9889 | -106.3918
13.9926 | -106.3522
13.9963 | -106.3141
14 | -106.2732
14.0037 | -106.2307
14.0074 | -106.1887
14.0111 | -106.148
14.0148 | -106.1083
14.0185 | -106.068
14.0222 | -106.0278
14.0259 | -105.9948
14.0296 | -105.9649
14.0333 | -105.9357
14.037 | -105.912
14.0407 | -105.8933
14.0444 | -105.8766
14.0481 | -105.8637
14.0519 | -105.8533
14.0556 | -105.8476
14.0593 | -105.8489
14.063 | -105.8532
14.0667 | -105.8619
14.0704 | -105.8718
14.0741 | -105.8857
14.0778 | -105.9018
14.0815 | -105.9214
14.0852 | -105.9427
14.0889 | -105.968
14.0926 | -105.9923
14.0963 | -106.0207
14.1 | -106.0502
14.1037 | -106.0822
14.1074 | -106.1142
14.1111 | -106.143
14.1148 | -106.1748
14.1185 | -106.2047
14.1222 | -106.2295
14.1259 | -106.2533
14.1296 | -106.2769
14.1333 | -106.302
14.137 | -106.329
14.1407 | -106.3512
14.1444 | -106.3717
14.1481 | -106.3932
14.1519 | -106.4122
14.1556 | -106.4308
14.1593 | -106.4519
14.163 | -106.4781
14.1667 | -106.5072
14.1704 | -106.5275
14.1741 | -106.5473
14.1778 | -106.5654
14.1815 | -106.5808
14.1852 | -106.5965
14.1889 | -106.613
14.1926 | -106.6311
14.1963 | -106.647
14.2 | -106.6614
14.2037 | -106.6777
14.2074 | -106.6949
14.2111 | -106.709
14.2148 | -106.7196
14.2185 | -106.7308
14.2222 | -106.7447
14.2259 | -106.7563
14.2296 | -106.7599
14.2333 | -106.7614
14.237 | -106.7615
14.2407 | -106.7586
14.2444 | -106.7535
14.2481 | -106.7498
14.2519 | -106.7454
14.2556 | -106.739
14.2593 | -106.7309
14.263 | -106.7211
14.2667 | -106.7125
14.2704 | -106.7021
14.2741 | -106.6963
14.2778 | -106.6888
14.2815 | -106.6795
14.2852 | -106.6703
14.2889 | -106.663
14.2926 | -106.6538
14.2963 | -106.6465
14.3 | -106.6409
14.3037 | -106.6327
14.3074 | -106.625
14.3111 | -106.617
14.3148 | -106.6114
14.3185 | -106.6052
14.3222 | -106.5993
14.3259 | -106.5968
14.3296 | -106.5959
14.3333 | -106.5945
14.337 | -106.5895
14.3407 | -106.5864
14.3444 | -106.5875
14.3481 | -106.5875
14.3519 | -106.5852
14.3556 | -106.5867
14.3593 | -106.5926
14.363 | -106.5957
14.3667 | -106.5999
14.3704 | -106.6043
14.3741 | -106.6088
14.3778 | -106.6136
14.3815 | -106.6188
14.3852 | -106.6227
14.3889 | -106.627
14.3926 | -106.6295
14.3963 | -106.6339
14.4 | -106.6412
14.4037 | -106.6474
14.4074 | -106.6537
14.4111 | -106.6575
14.4148 | -106.6617
14.4185 | -106.6647
14.4222 | -106.6665
14.4259 | -106.6594
14.4296 | -106.6584
14.4333 | -106.6528
14.437 | -106.6478
14.4407 | -106.6416
14.4444 | -106.6358
14.4481 | -106.6271
14.4519 | -106.6142
14.4556 | -106.5973
14.4593 | -106.5786
14.463 | -106.5641
14.4667 | -106.5471
14.4704 | -106.5263
14.4741 | -106.5043
14.4778 | -106.4837
14.4815 | -106.4611
14.4852 | -106.439
14.4889 | -106.4174
14.4926 | -106.394
14.4963 | -106.3699
14.5 | -106.3473
14.5037 | -106.3258
14.5074 | -106.3051
14.5111 | -106.2863
14.5148 | -106.2663
14.5185 | -106.2513
14.5222 | -106.2383
14.5259 | -106.2266
14.5296 | -106.2184
14.5333 | -106.2098
14.537 | -106.2007
14.5407 | -106.192
14.5444 | -106.1847
14.5481 | -106.1814
14.5519 | -106.1785
14.5556 | -106.181
14.5593 | -106.1885
14.563 | -106.1958
14.5667 | -106.2029
14.5704 | -106.2126
14.5741 | -106.2242
14.5778 | -106.2348
14.5815 | -106.2488
14.5852 | -106.2638
14.5889 | -106.2777
14.5926 | -106.2911
14.5963 | -106.3047
14.6 | -106.3214
14.6037 | -106.3344
14.6074 | -106.3446
14.6111 | -106.3547
14.6148 | -106.3652
14.6185 | -106.3764
14.6222 | -106.3857
14.6259 | -106.3941
14.6296 | -106.4037
14.6333 | -106.4169
14.637 | -106.4266
14.6407 | -106.4336
14.6444 | -106.4434
14.6481 | -106.4514
14.6519 | -106.4598
14.6556 | -106.4677
14.6593 | -106.4792
14.663 | -106.4912
14.6667 | -106.5008
14.6704 | -106.51
14.6741 | -106.5163
14.6778 | -106.5213
14.6815 | -106.5287
14.6852 | -106.5367
14.6889 | -106.5452
14.6926 | -106.5519
14.6963 | -106.5586
14.7 | -106.5672
14.7037 | -106.575
14.7074 | -106.5837
14.7111 | -106.5919
14.7148 | -106.6004
14.7185 | -106.612
14.7222 | -106.6247
14.7259 | -106.64
14.7296 | -106.6529
14.7333 | -106.6657
14.737 | -106.6785
14.7407 | -106.6888
14.7444 | -106.7015
14.7481 | -106.7171
14.7519 | -106.7337
14.7556 | -106.7543
14.7593 | -106.7709
14.763 | -106.7874
14.7667 | -106.8018
14.7704 | -106.8164
14.7741 | -106.8349
14.7778 | -106.8519
14.7815 | -106.8662
14.7852 | -106.8845
14.7889 | -106.9062
14.7926 | -106.9246
14.7963 | -106.9432
14.8 | -106.9619
14.8037 | -106.9822
14.8074 | -106.9996
14.8111 | -107.0167
14.8148 | -107.037
14.8185 | -107.0527
14.8222 | -107.0695
14.8259 | -107.0859
14.8296 | -107.1003
14.8333 | -107.1135
14.837 | -107.1243
14.8407 | -107.1278
14.8444 | -107.1288
14.8481 | -107.1299
14.8519 | -107.1281
14.8556 | -107.1234
14.8593 | -107.1189
14.863 | -107.1084
14.8667 | -107.0966
14.8704 | -107.0793
14.8741 | -107.0605
14.8778 | -107.0417
14.8815 | -107.019
14.8852 | -106.9945
14.8889 | -106.9677
14.8926 | -106.9416
14.8963 | -106.9127
14.9 | -106.883
14.9037 | -106.8528
14.9074 | -106.8218
14.9111 | -106.7916
14.9148 | -106.7619
14.9185 | -106.7307
14.9222 | -106.6982
14.9259 | -106.6636
14.9296 | -106.6298
14.9333 | -106.5994
14.937 | -106.5703
14.9407 | -106.541
14.9444 | -106.5122
14.9481 | -106.4867
14.9519 | -106.4659
14.9556 | -106.4472
14.9593 | -106.4372
14.963 | -106.4206
14.9667 | -106.4063
14.9704 | -106.3915
14.9741 | -106.3784
14.9778 | -106.3684
14.9815 | -106.3616
14.9852 | -106.3566
14.9889 | -106.3535
14.9926 | -106.3544
14.9963 | -106.3599
15 | -106.3643
15.0037 | -106.3675
15.0074 | -106.375
15.0111 | -106.3854
15.0148 | -106.3957
15.0185 | -106.4045
15.0222 | -106.4178
15.0259 | -106.431
15.0296 | -106.4448
15.0333 | -106.4642
15.037 | -106.4802
15.0407 | -106.4953
15.0444 | -106.511
15.0481 | -106.5278
15.0519 | -106.5465
15.0556 | -106.5631
15.0593 | -106.5772
15.063 | -106.5917
15.0667 | -106.6073
15.0704 | -106.6236
15.0741 | -106.6358
15.0778 | -106.6458
15.0815 | -106.6548
15.0852 | -106.6614
15.0889 | -106.6685
15.0926 | -106.6753
15.0963 | -106.6829
15.1 | -106.6878
15.1037 | -106.6925
15.1074 | -106.6926
15.1111 | -106.6963
15.1148 | -106.6997
15.1185 | -106.7044
15.1222 | -106.7071
15.1259 | -106.708
15.1296 | -106.7091
15.1333 | -106.7115
15.137 | -106.717
15.1407 | -106.7231
15.1444 | -106.7261
15.1481 | -106.7271
15.1519 | -106.7287
15.1556 | -106.7317
15.1593 | -106.7351
15.163 | -106.7371
15.1667 | -106.7373
15.1704 | -106.736
15.1741 | -106.7357
15.1778 | -106.7356
15.1815 | -106.7372
15.1852 | -106.736
15.1889 | -106.7308
15.1926 | -106.7285
15.1963 | -106.727
15.2 | -106.7244
15.2037 | -106.7211
15.2074 | -106.7154
15.2111 | -106.7089
15.2148 | -106.7035
15.2185 | -106.6972
15.2222 | -106.6937
15.2259 | -106.687
15.2296 | -106.6801
15.2333 | -106.675
15.237 | -106.6678
15.2407 | -106.659
15.2444 | -106.6504
15.2481 | -106.6396
15.2519 | -106.6283
15.2556 | -106.6194
15.2593 | -106.6125
15.263 | -106.6061
15.2667 | -106.6004
15.2704 | -106.5929
15.2741 | -106.5894
15.2778 | -106.5848
15.2815 | -106.5785
15.2852 | -106.5739
15.2889 | -106.5689
15.2926 | -106.5686
15.2963 | -106.5715
15.3 | -106.5729
15.3037 | -106.5763
15.3074 | -106.5795
15.3111 | -106.582
15.3148 | -106.5854
15.3185 | -106.5918
15.3222 | -106.597
15.3259 | -106.6005
15.3296 | -106.603
15.3333 | -106.6058
15.337 | -106.6096
15.3407 | -106.616
15.3444 | -106.6183
15.3481 | -106.6211
15.3519 | -106.6258
15.3556 | -106.6293
15.3593 | -106.6309
15.363 | -106.633
15.3667 | -106.6375
15.3704 | -106.6399
15.3741 | -106.6419
15.3778 | -106.6419
15.3815 | -106.6406
15.3852 | -106.6437
15.3889 | -106.6447
15.3926 | -106.6439
15.3963 | -106.6455
15.4 | -106.6446
15.4037 | -106.6431
15.4074 | -106.6449
15.4111 | -106.645
15.4148 | -106.6441
15.4185 | -106.6446
15.4222 | -106.6441
15.4259 | -106.6403
15.4296 | -106.637
15.4333 | -106.6331
15.437 | -106.6255
15.4407 | -106.6183
15.4444 | -106.6121
15.4481 | -106.6026
15.4519 | -106.5932
15.4556 | -106.5874
15.4593 | -106.5783
15.463 | -106.5661
15.4667 | -106.5538
15.4704 | -106.5405
15.4741 | -106.5264
15.4778 | -106.5101
15.4815 | -106.4948
15.4852 | -106.4808
15.4889 | -106.4662
15.4926 | -106.4545
15.4963 | -106.4422
15.5 | -106.4309
15.5037 | -106.4219
15.5074 | -106.4128
15.5111 | -106.4027
15.5148 | -106.3918
15.5185 | -106.3839
15.5222 | -106.38
15.5259 | -106.3796
15.5296 | -106.3792
15.5333 | -106.3787
15.537 | -106.3797
15.5407 | -106.3823
15.5444 | -106.3896
15.5481 | -106.4
15.5519 | -106.4092
15.5556 | -106.4208
15.5593 | -106.4342
15.563 | -106.449
15.5667 | -106.4659
15.5704 | -106.4838
15.5741 | -106.5003
15.5778 | -106.5184
15.5815 | -106.5409
15.5852 | -106.5615
15.5889 | -106.5819
15.5926 | -106.6035
15.5963 | -106.6214
15.6 | -106.6381
15.6037 | -106.6582
15.6074 | -106.6745
15.6111 | -106.6899
15.6148 | -106.7038
15.6185 | -106.7179
15.6222 | -106.7322
15.6259 | -106.7455
15.6296 | -106.754
15.6333 | -106.7624
15.637 | -106.7673
15.6407 | -106.7692
15.6444 | -106.7729
15.6481 | -106.7773
15.6519 | -106.7779
15.6556 | -106.7717
15.6593 | -106.7695
15.663 | -106.7653
15.6667 | -106.7588
15.6704 | -106.7524
15.6741 | -106.7445
15.6778 | -106.735
15.6815 | -106.725
15.6852 | -106.7157
15.6889 | -106.7058
15.6926 | -106.6949
15.6963 | -106.6846
15.7 | -106.6741
15.7037 | -106.6637
15.7074 | -106.6526
15.7111 | -106.6426
15.7148 | -106.6311
15.7185 | -106.6205
15.7222 | -106.6109
15.7259 | -106.6014
15.7296 | -106.5905
15.7333 | -106.5814
15.737 | -106.5742
15.7407 | -106.5676
15.7444 | -106.5599
15.7481 | -106.5529
15.7519 | -106.5473
15.7556 | -106.5417
15.7593 | -106.5374
15.763 | -106.5329
15.7667 | -106.5295
15.7704 | -106.5265
15.7741 | -106.5253
15.7778 | -106.5256
15.7815 | -106.5249
15.7852 | -106.5226
15.7889 | -106.5216
15.7926 | -106.523
15.7963 | -106.5244
15.8 | -106.5267
15.8037 | -106.5288
15.8074 | -106.5296
15.8111 | -106.5334
15.8148 | -106.5362
15.8185 | -106.5371
15.8222 | -106.5384
15.8259 | -106.5435
15.8296 | -106.5447
15.8333 | -106.5438
15.837 | -106.5461
15.8407 | -106.5477
15.8444 | -106.5421
15.8481 | -106.5322
15.8519 | -106.5263
15.8556 | -106.5201
15.8593 | -106.5106
15.863 | -106.5024
15.8667 | -106.4938
15.8704 | -106.4834
15.8741 | -106.4752
15.8778 | -106.4676
15.8815 | -106.461
15.8852 | -106.4499
15.8889 | -106.4367
15.8926 | -106.4242
15.8963 | -106.411
15.9 | -106.3996
15.9037 | -106.3871
15.9074 | -106.3754
15.9111 | -106.3675
15.9148 | -106.3607
15.9185 | -106.3528
15.9222 | -106.3429
15.9259 | -106.3374
15.9296 | -106.3344
15.9333 | -106.331
15.937 | -106.3298
15.9407 | -106.3293
15.9444 | -106.3305
15.9481 | -106.3332
15.9519 | -106.3389
15.9556 | -106.3452
15.9593 | -106.3503
15.963 | -106.3575
15.9667 | -106.3648
15.9704 | -106.3733
15.9741 | -106.3821
15.9778 | -106.3916
15.9815 | -106.4043
15.9852 | -106.4209
15.9889 | -106.4363
15.9926 | -106.4479
15.9963 | -106.4584
16 | -106.4696
16.0037 | -106.4824
16.0074 | -106.4956
16.0111 | -106.5096
16.0148 | -106.5221
16.0185 | -106.5342
16.0222 | -106.5452
16.0259 | -106.5553
16.0296 | -106.5619
16.0333 | -106.5708
16.037 | -106.578
16.0407 | -106.5841
16.0444 | -106.5926
16.0481 | -106.603
16.0519 | -106.6132
16.0556 | -106.6198
16.0593 | -106.6274
16.063 | -106.6358
16.0667 | -106.6404
16.0704 | -106.645
16.0741 | -106.6515
16.0778 | -106.6606
16.0815 | -106.6677
16.0852 | -106.6762
16.0889 | -106.6844
16.0926 | -106.6904
16.0963 | -106.6958
16.1 | -106.701
16.1037 | -106.7062
16.1074 | -106.7114
16.1111 | -106.7156
16.1148 | -106.7228
16.1185 | -106.7295
16.1222 | -106.7333
16.1259 | -106.7375
16.1296 | -106.7427
16.1333 | -106.7485
16.137 | -106.7542
16.1407 | -106.7588
16.1444 | -106.7626
16.1481 | -106.7668
16.1519 | -106.771
16.1556 | -106.7704
16.1593 | -106.768
16.163 | -106.7679
16.1667 | -106.7681
16.1704 | -106.7693
16.1741 | -106.7666
16.1778 | -106.765
16.1815 | -106.7641
16.1852 | -106.7614
16.1889 | -106.7609
16.1926 | -106.7561
16.1963 | -106.7519
16.2 | -106.7468
16.2037 | -106.7426
16.2074 | -106.7358
16.2111 | -106.7276
16.2148 | -106.7198
16.2185 | -106.7135
16.2222 | -106.7099
16.2259 | -106.706
16.2296 | -106.7009
16.2333 | -106.6947
16.237 | -106.6866
16.2407 | -106.6807
16.2444 | -106.6773
16.2481 | -106.6741
16.2519 | -106.671
16.2556 | -106.6669
16.2593 | -106.6636
16.263 | -106.6619
16.2667 | -106.6612
16.2704 | -106.6618
16.2741 | -106.6623
16.2778 | -106.6612
16.2815 | -106.6608
16.2852 | -106.6614
16.2889 | -106.6622
16.2926 | -106.6627
16.2963 | -106.6658
16.3 | -106.6705
16.3037 | -106.6755
16.3074 | -106.6821
16.3111 | -106.6872
16.3148 | -106.6891
16.3185 | -106.6907
16.3222 | -106.6942
16.3259 | -106.6994
16.3296 | -106.7049
16.3333 | -106.709
16.337 | -106.7121
16.3407 | -106.7165
16.3444 | -106.7221
16.3481 | -106.7258
16.3519 | -106.732
16.3556 | -106.7394
16.3593 | -106.7455
16.363 | -106.7486
16.3667 | -106.7508
16.3704 | -106.7513
16.3741 | -106.7508
16.3778 | -106.7478
16.3815 | -106.7473
16.3852 | -106.7492
16.3889 | -106.7515
16.3926 | -106.7513
16.3963 | -106.751
16.4 | -106.7507
16.4037 | -106.7476
16.4074 | -106.7441
16.4111 | -106.7402
16.4148 | -106.7339
16.4185 | -106.7301
16.4222 | -106.7249
16.4259 | -106.7245
16.4296 | -106.7226
16.4333 | -106.7193
16.437 | -106.7142
16.4407 | -106.7113
16.4444 | -106.7092
16.4481 | -106.7067
16.4519 | -106.7033
16.4556 | -106.7004
16.4593 | -106.6993
16.463 | -106.6981
16.4667 | -106.6969
16.4704 | -106.6958
16.4741 | -106.6894
16.4778 | -106.6863
16.4815 | -106.6825
16.4852 | -106.6793
16.4889 | -106.6766
16.4926 | -106.6767
16.4963 | -106.677
16.5 | -106.6767
16.5037 | -106.6765
16.5074 | -106.675
16.5111 | -106.6758
16.5148 | -106.6776
16.5185 | -106.6782
16.5222 | -106.6767
16.5259 | -106.6758
16.5296 | -106.6738
16.5333 | -106.6732
16.537 | -106.6729
16.5407 | -106.6705
16.5444 | -106.6715
16.5481 | -106.6758
16.5519 | -106.6776
16.5556 | -106.6777
16.5593 | -106.6769
16.563 | -106.6774
16.5667 | -106.6785
16.5704 | -106.6804
16.5741 | -106.6813
16.5778 | -106.6819
16.5815 | -106.6815
16.5852 | -106.6811
16.5889 | -106.6791
16.5926 | -106.6767
16.5963 | -106.6779
16.6 | -106.6783
16.6037 | -106.6793
16.6074 | -106.6828
16.6111 | -106.6855
16.6148 | -106.6904
16.6185 | -106.6948
16.6222 | -106.6983
16.6259 | -106.7007
16.6296 | -106.7033
16.6333 | -106.7052
16.637 | -106.7074
16.6407 | -106.7142
16.6444 | -106.721
16.6481 | -106.7277
16.6519 | -106.7342
16.6556 | -106.7426
16.6593 | -106.7519
16.663 | -106.758
16.6667 | -106.7652
16.6704 | -106.7775
16.6741 | -106.7907
16.6778 | -106.8021
16.6815 | -106.814
16.6852 | -106.8242
16.6889 | -106.8355
16.6926 | -106.8484
16.6963 | -106.8601
16.7 | -106.8726
16.7037 | -106.8876
16.7074 | -106.9048
16.7111 | -106.9246
16.7148 | -106.9392
16.7185 | -106.955
16.7222 | -106.9693
16.7259 | -106.9849
16.7296 | -107
16.7333 | -107.0131
16.737 | -107.0244
16.7407 | -107.0359
16.7444 | -107.05
16.7481 | -107.0605
16.7519 | -107.0699
16.7556 | -107.0812
16.7593 | -107.0913
16.763 | -107.1004
16.7667 | -107.1065
16.7704 | -107.1119
16.7741 | -107.1169
16.7778 | -107.1202
16.7815 | -107.1157
16.7852 | -107.1091
16.7889 | -107.1057
16.7926 | -107.1012
16.7963 | -107.0971
16.8 | -107.0909
16.8037 | -107.081
16.8074 | -107.0671
16.8111 | -107.0536
16.8148 | -107.0402
16.8185 | -107.0246
16.8222 | -107.0092
16.8259 | -106.9945
16.8296 | -106.981
16.8333 | -106.9693
16.837 | -106.9526
16.8407 | -106.9341
16.8444 | -106.9146
16.8481 | -106.8963
16.8519 | -106.8813
16.8556 | -106.8681
16.8593 | -106.8541
16.863 | -106.8398
16.8667 | -106.8276
16.8704 | -106.8171
16.8741 | -106.8078
16.8778 | -106.7989
16.8815 | -106.7872
16.8852 | -106.7796
16.8889 | -106.7724
16.8926 | -106.7675
16.8963 | -106.7626
16.9 | -106.7591
16.9037 | -106.7582
16.9074 | -106.7579
16.9111 | -106.7552
16.9148 | -106.7539
16.9185 | -106.7545
16.9222 | -106.7548
16.9259 | -106.7579
16.9296 | -106.7638
16.9333 | -106.7697
16.937 | -106.7774
16.9407 | -106.7853
16.9444 | -106.79
16.9481 | -106.7959
16.9519 | -106.8022
16.9556 | -106.809
16.9593 | -106.8146
16.963 | -106.8226
16.9667 | -106.8307
16.9704 | -106.8372
16.9741 | -106.8467
16.9778 | -106.8577
16.9815 | -106.8691
16.9852 | -106.8786
16.9889 | -106.8866
16.9926 | -106.8928
16.9963 | -106.8997
17 | -106.9068
17.0037 | -106.9129
17.0074 | -106.9204
17.0111 | -106.9272
17.0148 | -106.9325
17.0185 | -106.9376
17.0222 | -106.9458
17.0259 | -106.9524
17.0296 | -106.954
17.0333 | -106.9559
17.037 | -106.9575
17.0407 | -106.9574
17.0444 | -106.9607
17.0481 | -106.9638
17.0519 | -106.964
17.0556 | -106.9644
17.0593 | -106.965
17.063 | -106.9661
17.0667 | -106.9666
17.0704 | -106.9651
17.0741 | -106.9618
17.0778 | -106.9613
17.0815 | -106.9648
17.0852 | -106.9679
17.0889 | -106.9687
17.0926 | -106.968
17.0963 | -106.9721
17.1 | -106.9742
17.1037 | -106.9779
17.1074 | -106.9807
17.1111 | -106.9836
17.1148 | -106.9897
17.1185 | -106.9938
17.1222 | -106.999
17.1259 | -107.0057
17.1296 | -107.0099
17.1333 | -107.0138
17.137 | -107.0186
17.1407 | -107.0209
17.1444 | -107.0227
17.1481 | -107.0243
17.1519 | -107.025
17.1556 | -107.0272
17.1593 | -107.0312
17.163 | -107.0344
17.1667 | -107.0373
17.1704 | -107.0371
17.1741 | -107.0328
17.1778 | -107.0282
17.1815 | -107.0241
17.1852 | -107.0204
17.1889 | -107.019
17.1926 | -107.0156
17.1963 | -107.0112
17.2 | -107.0075
17.2037 | -107.0029
17.2074 | -106.9946
17.2111 | -106.9858
17.2148 | -106.9761
17.2185 | -106.9656
17.2222 | -106.9542
17.2259 | -106.9442
17.2296 | -106.9348
17.2333 | -106.9265
17.237 | -106.9194
17.2407 | -106.9099
17.2444 | -106.9011
17.2481 | -106.8903
17.2519 | -106.8788
17.2556 | -106.8663
17.2593 | -106.8562
17.263 | -106.8492
17.2667 | -106.8408
17.2704 | -106.835
17.2741 | -106.8314
17.2778 | -106.8255
17.2815 | -106.8212
17.2852 | -106.8181
17.2889 | -106.8127
17.2926 | -106.8122
17.2963 | -106.8142
17.3 | -106.8144
17.3037 | -106.8143
17.3074 | -106.8153
17.3111 | -106.817
17.3148 | -106.8211
17.3185 | -106.8282
17.3222 | -106.8349
17.3259 | -106.843
17.3296 | -106.8537
17.3333 | -106.8634
17.337 | -106.8718
17.3407 | -106.8797
17.3444 | -106.889
17.3481 | -106.8982
17.3519 | -106.9079
17.3556 | -106.9192
17.3593 | -106.9287
17.363 | -106.9357
17.3667 | -106.9458
17.3704 | -106.9549
17.3741 | -106.9631
17.3778 | -106.9723
17.3815 | -106.9824
17.3852 | -106.9891
17.3889 | -106.9966
17.3926 | -107.0018
17.3963 | -107.0045
17.4 | -107.0085
17.4037 | -107.0114
17.4074 | -107.0132
17.4111 | -107.0159
17.4148 | -107.0184
17.4185 | -107.0202
17.4222 | -107.0185
17.4259 | -107.0185
17.4296 | -107.0202
17.4333 | -107.0183
17.437 | -107.0127
17.4407 | -107.0084
17.4444 | -107.004
17.4481 | -107
17.4519 | -106.9978
17.4556 | -106.9952
17.4593 | -106.9916
17.463 | -106.9879
17.4667 | -106.9828
17.4704 | -106.9781
17.4741 | -106.9719
17.4778 | -106.9659
17.4815 | -106.9611
17.4852 | -106.9569
17.4889 | -106.9496
17.4926 | -106.9404
17.4963 | -106.9341
17.5 | -106.9264
17.5037 | -106.92
17.5074 | -106.911
17.5111 | -106.8979
17.5148 | -106.8879
17.5185 | -106.8775
17.5222 | -106.8667
17.5259 | -106.8591
17.5296 | -106.8524
17.5333 | -106.8448
17.537 | -106.8384
17.5407 | -106.8304
17.5444 | -106.8209
17.5481 | -106.8123
17.5519 | -106.8062
17.5556 | -106.7998
17.5593 | -106.796
17.563 | -106.7929
17.5667 | -106.789
17.5704 | -106.7875
17.5741 | -106.7859
17.5778 | -106.7829
17.5815 | -106.7831
17.5852 | -106.787
17.5889 | -106.7919
17.5926 | -106.7964
17.5963 | -106.8004
17.6 | -106.8066
17.6037 | -106.815
17.6074 | -106.8259
17.6111 | -106.8346
17.6148 | -106.8432
17.6185 | -106.851
17.6222 | -106.8576
17.6259 | -106.8649
17.6296 | -106.8739
17.6333 | -106.881
17.637 | -106.89
17.6407 | -106.8987
17.6444 | -106.9094
17.6481 | -106.9194
17.6519 | -106.9276
17.6556 | -106.9354
17.6593 | -106.9414
17.663 | -106.9477
17.6667 | -106.9531
17.6704 | -106.9595
17.6741 | -106.9669
17.6778 | -106.972
17.6815 | -106.9729
17.6852 | -106.9754
17.6889 | -106.9791
17.6926 | -106.9807
17.6963 | -106.9815
17.7 | -106.9799
17.7037 | -106.9761
17.7074 | -106.9743
17.7111 | -106.9713
17.7148 | -106.9659
17.7185 | -106.9608
17.7222 | -106.9544
17.7259 | -106.9491
17.7296 | -106.9428
17.7333 | -106.9365
17.737 | -106.9289
17.7407 | -106.9233
17.7444 | -106.9194
17.7481 | -106.9128
17.7519 | -106.9058
17.7556 | -106.897
17.7593 | -106.8875
17.763 | -106.8824
17.7667 | -106.8745
17.7704 | -106.866
17.7741 | -106.859
17.7778 | -106.8508
17.7815 | -106.8442
17.7852 | -106.8379
17.7889 | -106.8328
17.7926 | -106.8284
17.7963 | -106.8238
17.8 | -106.8212
17.8037 | -106.8194
17.8074 | -106.8192
17.8111 | -106.817
17.8148 | -106.8164
17.8185 | -106.8178
17.8222 | -106.8163
17.8259 | -106.8171
17.8296 | -106.8223
17.8333 | -106.8251
17.837 | -106.8295
17.8407 | -106.8337
17.8444 | -106.8371
17.8481 | -106.8432
17.8519 | -106.8462
17.8556 | -106.8512
17.8593 | -106.8588
17.863 | -106.8688
17.8667 | -106.8771
17.8704 | -106.887
17.8741 | -106.8991
17.8778 | -106.9087
17.8815 | -106.9163
17.8852 | -106.9261
17.8889 | -106.9348
17.8926 | -106.9428
17.8963 | -106.9498
17.9 | -106.9557
17.9037 | -106.9649
17.9074 | -106.9731
17.9111 | -106.9787
17.9148 | -106.9835
17.9185 | -106.9883
17.9222 | -106.9917
17.9259 | -106.9945
17.9296 | -106.9995
17.9333 | -106.999
17.937 | -106.9992
17.9407 | -107.0013
17.9444 | -107.0037
17.9481 | -107.0055
17.9519 | -107.0051
17.9556 | -107.0043
17.9593 | -107.0046
17.963 | -107.0028
17.9667 | -107.0014
17.9704 | -107.0023
17.9741 | -107.0034
17.9778 | -107.0036
17.9815 | -107.0035
17.9852 | -107.005
17.9889 | -107.0044
17.9926 | -107.003
17.9963 | -107.0036
18 | -107.0035
18.0037 | -107.0049
18.0074 | -107.0063
18.0111 | -107.0072
18.0148 | -107.0091
18.0185 | -107.0089
18.0222 | -107.0084
18.0259 | -107.0048
18.0296 | -107.0012
18.0333 | -106.9982
18.037 | -106.992
18.0407 | -106.988
18.0444 | -106.9843
18.0481 | -106.98
18.0519 | -106.9759
18.0556 | -106.9706
18.0593 | -106.9637
18.063 | -106.9586
18.0667 | -106.9536
18.0704 | -106.9485
18.0741 | -106.9428
18.0778 | -106.9394
18.0815 | -106.9342
18.0852 | -106.9262
18.0889 | -106.922
18.0926 | -106.9176
18.0963 | -106.9153
18.1 | -106.9145
18.1037 | -106.9122
18.1074 | -106.9126
18.1111 | -106.9116
18.1148 | -106.9097
18.1185 | -106.909
18.1222 | -106.9092
18.1259 | -106.9084
18.1296 | -106.908
18.1333 | -106.9091
18.137 | -106.9109
18.1407 | -106.9107
18.1444 | -106.9126
18.1481 | -106.9151
18.1519 | -106.9183
18.1556 | -106.9212
18.1593 | -106.9207
18.163 | -106.9193
18.1667 | -106.918
18.1704 | -106.9198
18.1741 | -106.9207
18.1778 | -106.9217
18.1815 | -106.924
18.1852 | -106.9238
18.1889 | -106.9209
18.1926 | -106.9209
18.1963 | -106.9206
18.2 | -106.9187
18.2037 | -106.9155
18.2074 | -106.911
18.2111 | -106.9067
18.2148 | -106.9023
18.2185 | -106.8978
18.2222 | -106.8928
18.2259 | -106.8842
18.2296 | -106.8758
18.2333 | -106.8684
18.237 | -106.863
18.2407 | -106.8555
18.2444 | -106.8479
18.2481 | -106.8414
18.2519 | -106.8355
18.2556 | -106.8287
18.2593 | -106.8202
18.263 | -106.8118
18.2667 | -106.8055
18.2704 | -106.7998
18.2741 | -106.7942
18.2778 | -106.7892
18.2815 | -106.7852
18.2852 | -106.784
18.2889 | -106.7815
18.2926 | -106.7787
18.2963 | -106.7775
18.3 | -106.7755
18.3037 | -106.7756
18.3074 | -106.777
18.3111 | -106.7806
18.3148 | -106.7838
18.3185 | -106.7887
18.3222 | -106.7927
18.3259 | -106.8
18.3296 | -106.8074
18.3333 | -106.8169
18.337 | -106.826
18.3407 | -106.8348
18.3444 | -106.8458
18.3481 | -106.8559
18.3519 | -106.8656
18.3556 | -106.8737
18.3593 | -106.882
18.363 | -106.8931
18.3667 | -106.9046
18.3704 | -106.9151
18.3741 | -106.9264
18.3778 | -106.9383
18.3815 | -106.9486
18.3852 | -106.9581
18.3889 | -106.9674
18.3926 | -106.9763
18.3963 | -106.9848
18.4 | -106.9915
18.4037 | -106.9992
18.4074 | -107.0065
18.4111 | -107.0121
18.4148 | -107.0164
18.4185 | -107.0219
18.4222 | -107.0291
18.4259 | -107.0324
18.4296 | -107.0352
18.4333 | -107.0357
18.437 | -107.0396
18.4407 | -107.042
18.4444 | -107.0435
18.4481 | -107.047
18.4519 | -107.051
18.4556 | -107.0517
18.4593 | -107.0504
18.463 | -107.0486
18.4667 | -107.0481
18.4704 | -107.0499
18.4741 | -107.051
18.4778 | -107.0498
18.4815 | -107.0493
18.4852 | -107.0508
18.4889 | -107.0548
18.4926 | -107.058
18.4963 | -107.0583
18.5 | -107.0578
18.5037 | -107.0589
18.5074 | -107.0605
18.5111 | -107.0642
18.5148 | -107.069
18.5185 | -107.0712
18.5222 | -107.0746
18.5259 | -107.0783
18.5296 | -107.0823
18.5333 | -107.0839
18.537 | -107.0856
18.5407 | -107.0852
18.5444 | -107.0868
18.5481 | -107.0891
18.5519 | -107.0932
18.5556 | -107.0943
18.5593 | -107.0935
18.563 | -107.0935
18.5667 | -107.0915
18.5704 | -107.0867
18.5741 | -107.0827
18.5778 | -107.0795
18.5815 | -107.075
18.5852 | -107.0703
18.5889 | -107.0634
18.5926 | -107.0572
18.5963 | -107.0503
18.6 | -107.0477
18.6037 | -107.0442
18.6074 | -107.0377
18.6111 | -107.0301
18.6148 | -107.0227
18.6185 | -107.0187
18.6222 | -107.0156
18.6259 | -107.0106
18.6296 | -107.0054
18.6333 | -107.0009
18.637 | -106.9979
18.6407 | -106.9953
18.6444 | -106.9936
18.6481 | -106.9934
18.6519 | -106.9941
18.6556 | -106.9964
18.6593 | -106.9993
18.663 | -107.0014
18.6667 | -107.0007
18.6704 | -107.0019
18.6741 | -107.0066
18.6778 | -107.013
18.6815 | -107.0172
18.6852 | -107.0206
18.6889 | -107.0267
18.6926 | -107.0337
18.6963 | -107.0408
18.7 | -107.0529
18.7037 | -107.0579
18.7074 | -107.0642
18.7111 | -107.0728
18.7148 | -107.0805
18.7185 | -107.0884
18.7222 | -107.0993
18.7259 | -107.1105
18.7296 | -107.1217
18.7333 | -107.1314
18.737 | -107.1409
18.7407 | -107.1524
18.7444 | -107.1652
18.7481 | -107.1742
18.7519 | -107.1833
18.7556 | -107.1931
18.7593 | -107.1997
18.763 | -107.2095
18.7667 | -107.2207
18.7704 | -107.2305
18.7741 | -107.2393
18.7778 | -107.2469
18.7815 | -107.2541
18.7852 | -107.2643
18.7889 | -107.2746
18.7926 | -107.2831
18.7963 | -107.2906
18.8 | -107.2997
18.8037 | -107.3062
18.8074 | -107.313
18.8111 | -107.3157
18.8148 | -107.3162
18.8185 | -107.3172
18.8222 | -107.3206
18.8259 | -107.3243
18.8296 | -107.3272
18.8333 | -107.3269
18.837 | -107.3285
18.8407 | -107.3292
18.8444 | -107.3289
18.8481 | -107.3273
18.8519 | -107.3243
18.8556 | -107.3232
18.8593 | -107.3212
18.863 | -107.3176
18.8667 | -107.3116
18.8704 | -107.306
18.8741 | -107.302
18.8778 | -107.2976
18.8815 | -107.2938
18.8852 | -107.2871
18.8889 | -107.2784
18.8926 | -107.2725
18.8963 | -107.266
18.9 | -107.2599
18.9037 | -107.2532
18.9074 | -107.2464
18.9111 | -107.2411
18.9148 | -107.2361
18.9185 | -107.2297
18.9222 | -107.2212
18.9259 | -107.2124
18.9296 | -107.2048
18.9333 | -107.1993
18.937 | -107.1936
18.9407 | -107.1879
18.9444 | -107.1777
18.9481 | -107.1691
18.9519 | -107.1622
18.9556 | -107.1539
18.9593 | -107.1436
18.963 | -107.1321
18.9667 | -107.1191
18.9704 | -107.11
18.9741 | -107.1029
18.9778 | -107.0945
18.9815 | -107.0853
18.9852 | -107.0781
18.9889 | -107.0702
18.9926 | -107.0621
18.9963 | -107.0547
19 | -107.0471
19.0037 | -107.0396
19.0074 | -107.031
19.0111 | -107.0215
19.0148 | -107.013
19.0185 | -107.0064
19.0222 | -106.9993
19.0259 | -106.9908
19.0296 | -106.9863
19.0333 | -106.979
19.037 | -106.9688
19.0407 | -106.9598
19.0444 | -106.9519
19.0481 | -106.9452
19.0519 | -106.9405
19.0556 | -106.9348
19.0593 | -106.9281
19.063 | -106.9224
19.0667 | -106.9177
19.0704 | -106.9141
19.0741 | -106.9136
19.0778 | -106.9128
19.0815 | -106.9121
19.0852 | -106.9127
19.0889 | -106.911
19.0926 | -106.9088
19.0963 | -106.9068
19.1 | -106.9066
19.1037 | -106.908
19.1074 | -106.9067
19.1111 | -106.906
19.1148 | -106.906
19.1185 | -106.9072
19.1222 | -106.9091
19.1259 | -106.9097
19.1296 | -106.9108
19.1333 | -106.9127
19.137 | -106.9134
19.1407 | -106.9149
19.1444 | -106.9163
19.1481 | -106.9184
19.1519 | -106.9196
19.1556 | -106.9225
19.1593 | -106.9273
19.163 | -106.9313
19.1667 | -106.9361
19.1704 | -106.9392
19.1741 | -106.9423
19.1778 | -106.9467
19.1815 | -106.9518
19.1852 | -106.9581
19.1889 | -106.9662
19.1926 | -106.9752
19.1963 | -106.9857
19.2 | -106.9973
19.2037 | -107.0069
19.2074 | -107.0169
19.2111 | -107.0264
19.2148 | -107.0402
19.2185 | -107.0549
19.2222 | -107.0664
19.2259 | -107.0797
19.2296 | -107.0953
19.2333 | -107.1089
19.237 | -107.1221
19.2407 | -107.1363
19.2444 | -107.1489
19.2481 | -107.1601
19.2519 | -107.1708
19.2556 | -107.1823
19.2593 | -107.1939
19.263 | -107.2056
19.2667 | -107.2163
19.2704 | -107.2281
19.2741 | -107.2398
19.2778 | -107.2503
19.2815 | -107.2603
19.2852 | -107.27
19.2889 | -107.2795
19.2926 | -107.2895
19.2963 | -107.2982
19.3 | -107.3065
19.3037 | -107.3125
19.3074 | -107.3186
19.3111 | -107.3262
19.3148 | -107.3333
19.3185 | -107.3372
19.3222 | -107.339
19.3259 | -107.3402
19.3296 | -107.3416
19.3333 | -107.3419
19.337 | -107.3415
19.3407 | -107.3402
19.3444 | -107.3378
19.3481 | -107.3339
19.3519 | -107.328
19.3556 | -107.3219
19.3593 | -107.3157
19.363 | -107.3105
19.3667 | -107.3033
19.3704 | -107.2963
19.3741 | -107.2858
19.3778 | -107.2756
19.3815 | -107.2672
19.3852 | -107.2573
19.3889 | -107.2475
19.3926 | -107.2375
19.3963 | -107.2263
19.4 | -107.2123
19.4037 | -107.1984
19.4074 | -107.1881
19.4111 | -107.1782
19.4148 | -107.1683
19.4185 | -107.1599
19.4222 | -107.1499
19.4259 | -107.1396
19.4296 | -107.1294
19.4333 | -107.1201
19.437 | -107.1119
19.4407 | -107.1036
19.4444 | -107.0993
19.4481 | -107.0937
19.4519 | -107.0882
19.4556 | -107.0845
19.4593 | -107.0806
19.463 | -107.0764
19.4667 | -107.0737
19.4704 | -107.0711
19.4741 | -107.0706
19.4778 | -107.072
19.4815 | -107.0749
19.4852 | -107.0773
19.4889 | -107.0804
19.4926 | -107.0833
19.4963 | -107.0869
19.5 | -107.0891
19.5037 | -107.0902
19.5074 | -107.0949
19.5111 | -107.1009
19.5148 | -107.1073
19.5185 | -107.1106
19.5222 | -107.1119
19.5259 | -107.116
19.5296 | -107.1218
19.5333 | -107.1252
19.537 | -107.1277
19.5407 | -107.1324
19.5444 | -107.1374
19.5481 | -107.1409
19.5519 | -107.1429
19.5556 | -107.1456
19.5593 | -107.1481
19.563 | -107.1503
19.5667 | -107.1505
19.5704 | -107.1513
19.5741 | -107.152
19.5778 | -107.154
19.5815 | -107.1544
19.5852 | -107.1556
19.5889 | -107.1544
19.5926 | -107.1517
19.5963 | -107.1506
19.6 | -107.1499
19.6037 | -107.1459
19.6074 | -107.1392
19.6111 | -107.1327
19.6148 | -107.1274
19.6185 | -107.1215
19.6222 | -107.1162
19.6259 | -107.1085
19.6296 | -107.1015
19.6333 | -107.0926
19.637 | -107.0823
19.6407 | -107.0717
19.6444 | -107.0608
19.6481 | -107.0476
19.6519 | -107.034
19.6556 | -107.0205
19.6593 | -107.0068
19.663 | -106.9928
19.6667 | -106.9772
19.6704 | -106.9615
19.6741 | -106.9467
19.6778 | -106.9348
19.6815 | -106.9225
19.6852 | -106.906
19.6889 | -106.8922
19.6926 | -106.8805
19.6963 | -106.8688
19.7 | -106.8594
19.7037 | -106.85
19.7074 | -106.8424
19.7111 | -106.8351
19.7148 | -106.8288
19.7185 | -106.8253
19.7222 | -106.8218
19.7259 | -106.8202
19.7296 | -106.8176
19.7333 | -106.816
19.737 | -106.8165
19.7407 | -106.8174
19.7444 | -106.82
19.7481 | -106.8238
19.7519 | -106.8266
19.7556 | -106.8305
19.7593 | -106.8355
19.763 | -106.842
19.7667 | -106.8565
19.7704 | -106.8648
19.7741 | -106.874
19.7778 | -106.8839
19.7815 | -106.8916
19.7852 | -106.9024
19.7889 | -106.9126
19.7926 | -106.9243
19.7963 | -106.9369
19.8 | -106.9498
19.8037 | -106.9619
19.8074 | -106.972
19.8111 | -106.9831
19.8148 | -106.9965
19.8185 | -107.008
19.8222 | -107.0192
19.8259 | -107.03
19.8296 | -107.0415
19.8333 | -107.0526
19.837 | -107.0654
19.8407 | -107.0766
19.8444 | -107.0854
19.8481 | -107.0953
19.8519 | -107.1024
19.8556 | -107.1068
19.8593 | -107.111
19.863 | -107.1161
19.8667 | -107.1185
19.8704 | -107.1204
19.8741 | -107.1214
19.8778 | -107.1193
19.8815 | -107.1165
19.8852 | -107.1145
19.8889 | -107.1106
19.8926 | -107.1087
19.8963 | -107.1036
19.9 | -107.0947
19.9037 | -107.0837
19.9074 | -107.0741
19.9111 | -107.0668
19.9148 | -107.0601
19.9185 | -107.0518
19.9222 | -107.0416
19.9259 | -107.0317
19.9296 | -107.023
19.9333 | -107.0161
19.937 | -107.0079
19.9407 | -106.9986
19.9444 | -106.9913
19.9481 | -106.9861
19.9519 | -106.981
19.9556 | -106.9762
19.9593 | -106.9715
19.963 | -106.9675
19.9667 | -106.9638
19.9704 | -106.9591
19.9741 | -106.9543
19.9778 | -106.9511
19.9815 | -106.9491
19.9852 | -106.9484
19.9889 | -106.9476
19.9926 | -106.9474
19.9963 | -106.9475
20 | -106.9495
20.0037 | -106.9521
20.0074 | -106.9555
20.0111 | -106.963
20.0148 | -106.9698
20.0185 | -106.9747
20.0222 | -106.9819
20.0259 | -106.9877
20.0296 | -106.9934
20.0333 | -107.0016
20.037 | -107.0099
20.0407 | -107.0163
20.0444 | -107.0254
20.0481 | -107.0343
20.0519 | -107.0393
20.0556 | -107.0443
20.0593 | -107.0507
20.063 | -107.0591
20.0667 | -107.0696
20.0704 | -107.0774
20.0741 | -107.0826
20.0778 | -107.0841
20.0815 | -107.087
20.0852 | -107.0892
20.0889 | -107.0925
20.0926 | -107.0955
20.0963 | -107.0965
20.1 | -107.0964
20.1037 | -107.0982
20.1074 | -107.0974
20.1111 | -107.095
20.1148 | -107.0913
20.1185 | -107.0874
20.1222 | -107.0824
20.1259 | -107.0772
20.1296 | -107.0711
20.1333 | -107.0632
20.137 | -107.0554
20.1407 | -107.05
20.1444 | -107.0438
20.1481 | -107.0378
20.1519 | -107.0323
20.1556 | -107.0258
20.1593 | -107.0193
20.163 | -107.0123
20.1667 | -107.0011
20.1704 | -106.9911
20.1741 | -106.9844
20.1778 | -106.9774
20.1815 | -106.9702
20.1852 | -106.9654
20.1889 | -106.9622
20.1926 | -106.9558
20.1963 | -106.948
20.2 | -106.9423
20.2037 | -106.9367
20.2074 | -106.9312
20.2111 | -106.9258
20.2148 | -106.9194
20.2185 | -106.9128
20.2222 | -106.9061
20.2259 | -106.9017
20.2296 | -106.8972
20.2333 | -106.8924
20.237 | -106.8868
20.2407 | -106.8796
20.2444 | -106.8744
20.2481 | -106.8705
20.2519 | -106.8678
20.2556 | -106.8659
20.2593 | -106.8633
20.263 | -106.8602
20.2667 | -106.8569
20.2704 | -106.8557
20.2741 | -106.8554
20.2778 | -106.854
20.2815 | -106.8534
20.2852 | -106.8527
20.2889 | -106.8514
20.2926 | -106.8521
20.2963 | -106.8524
20.3 | -106.8543
20.3037 | -106.859
20.3074 | -106.8627
20.3111 | -106.8643
20.3148 | -106.8679
20.3185 | -106.8724
20.3222 | -106.875
20.3259 | -106.8779
20.3296 | -106.8805
20.3333 | -106.8827
20.337 | -106.8856
20.3407 | -106.8901
20.3444 | -106.8968
20.3481 | -106.9024
20.3519 | -106.9069
20.3556 | -106.9114
20.3593 | -106.9145
20.363 | -106.9177
20.3667 | -106.9234
20.3704 | -106.9286
20.3741 | -106.9349
20.3778 | -106.94
20.3815 | -106.9442
20.3852 | -106.9485
20.3889 | -106.9525
20.3926 | -106.9557
20.3963 | -106.9593
20.4 | -106.9633
20.4037 | -106.9673
20.4074 | -106.9716
20.4111 | -106.9724
20.4148 | -106.9726
20.4185 | -106.9747
20.4222 | -106.9771
20.4259 | -106.9774
20.4296 | -106.9758
20.4333 | -106.9735
20.437 | -106.9706
20.4407 | -106.9671
20.4444 | -106.9637
20.4481 | -106.9602
20.4519 | -106.9573
20.4556 | -106.9568
20.4593 | -106.9548
20.463 | -106.9504
20.4667 | -106.9442
20.4704 | -106.9383
20.4741 | -106.9334
20.4778 | -106.9258
20.4815 | -106.9204
20.4852 | -106.9148
20.4889 | -106.9098
20.4926 | -106.9051
20.4963 | -106.9006
20.5 | -106.8967
20.5037 | -106.8926
20.5074 | -106.8865
20.5111 | -106.8823
20.5148 | -106.8778
20.5185 | -106.8722
20.5222 | -106.8672
20.5259 | -106.8628
20.5296 | -106.86
20.5333 | -106.855
20.537 | -106.8492
20.5407 | -106.8444
20.5444 | -106.8408
20.5481 | -106.8382
20.5519 | -106.8363
20.5556 | -106.8336
20.5593 | -106.831
20.563 | -106.8281
20.5667 | -106.827
20.5704 | -106.826
20.5741 | -106.8253
20.5778 | -106.8263
20.5815 | -106.8267
20.5852 | -106.8265
20.5889 | -106.8267
20.5926 | -106.8284
20.5963 | -106.831
20.6 | -106.8328
20.6037 | -106.8362
20.6074 | -106.8377
20.6111 | -106.839
20.6148 | -106.8411
20.6185 | -106.8445
20.6222 | -106.8488
20.6259 | -106.8519
20.6296 | -106.856
20.6333 | -106.8586
20.637 | -106.8586
20.6407 | -106.8614
20.6444 | -106.8638
20.6481 | -106.8653
20.6519 | -106.8664
20.6556 | -106.8671
20.6593 | -106.8648
20.663 | -106.8632
20.6667 | -106.8612
20.6704 | -106.8608
20.6741 | -106.8603
20.6778 | -106.8586
20.6815 | -106.8564
20.6852 | -106.8542
20.6889 | -106.8523
20.6926 | -106.8498
20.6963 | -106.8455
20.7 | -106.8429
20.7037 | -106.8412
20.7074 | -106.8385
20.7111 | -106.8337
20.7148 | -106.8257
20.7185 | -106.8193
20.7222 | -106.8131
20.7259 | -106.8063
20.7296 | -106.8001
20.7333 | -106.7943
20.737 | -106.7889
20.7407 | -106.7821
20.7444 | -106.7752
20.7481 | -106.7674
20.7519 | -106.7588
20.7556 | -106.751
20.7593 | -106.7436
20.763 | -106.7371
20.7667 | -106.7283
20.7704 | -106.7211
20.7741 | -106.716
20.7778 | -106.7104
20.7815 | -106.7041
20.7852 | -106.7003
20.7889 | -106.6973
20.7926 | -106.6949
20.7963 | -106.693
20.8 | -106.692
20.8037 | -106.6913
20.8074 | -106.6903
20.8111 | -106.6873
20.8148 | -106.6848
20.8185 | -106.6851
20.8222 | -106.6851
20.8259 | -106.6862
20.8296 | -106.6873
20.8333 | -106.6879
20.837 | -106.6908
20.8407 | -106.6899
20.8444 | -106.6891
20.8481 | -106.6883
20.8519 | -106.6893
20.8556 | -106.6888
20.8593 | -106.6886
20.863 | -106.69
20.8667 | -106.6907
20.8704 | -106.6892
20.8741 | -106.6873
20.8778 | -106.6853
20.8815 | -106.6853
20.8852 | -106.6831
20.8889 | -106.6824
20.8926 | -106.6806
20.8963 | -106.6788
20.9 | -106.6786
20.9037 | -106.6788
20.9074 | -106.6791
20.9111 | -106.6803
20.9148 | -106.6803
20.9185 | -106.6805
20.9222 | -106.6811
20.9259 | -106.6801
20.9296 | -106.6816
20.9333 | -106.6837
20.937 | -106.6822
20.9407 | -106.6817
20.9444 | -106.6849
20.9481 | -106.6896
20.9519 | -106.6939
20.9556 | -106.6998
20.9593 | -106.7043
20.963 | -106.7095
20.9667 | -106.7149
20.9704 | -106.7205
20.9741 | -106.7254
20.9778 | -106.733
20.9815 | -106.7417
20.9852 | -106.7497
20.9889 | -106.7583
20.9926 | -106.7683
20.9963 | -106.7776
21 | -106.7872
21.0037 | -106.7997
21.0074 | -106.8136
21.0111 | -106.8282
21.0148 | -106.8424
21.0185 | -106.8572
21.0222 | -106.8701
21.0259 | -106.8801
21.0296 | -106.8922
21.0333 | -106.9049
21.037 | -106.9191
21.0407 | -106.933
21.0444 | -106.9459
21.0481 | -106.9585
21.0519 | -106.9709
21.0556 | -106.9809
21.0593 | -106.9897
21.063 | -107.0023
21.0667 | -107.015
21.0704 | -107.0244
21.0741 | -107.0328
21.0778 | -107.0384
21.0815 | -107.0425
21.0852 | -107.0465
21.0889 | -107.0475
21.0926 | -107.0492
21.0963 | -107.0494
21.1 | -107.0462
21.1037 | -107.0435
21.1074 | -107.0439
21.1111 | -107.0422
21.1148 | -107.0388
21.1185 | -107.0332
21.1222 | -107.0251
21.1259 | -107.0151
21.1296 | -107.0044
21.1333 | -106.9932
21.137 | -106.9818
21.1407 | -106.9711
21.1444 | -106.9593
21.1481 | -106.9486
21.1519 | -106.9354
21.1556 | -106.9217
21.1593 | -106.9084
21.163 | -106.8938
21.1667 | -106.8808
21.1704 | -106.8704
21.1741 | -106.8597
21.1778 | -106.8488
21.1815 | -106.8341
21.1852 | -106.8224
21.1889 | -106.8114
21.1926 | -106.8011
21.1963 | -106.7914
21.2 | -106.7807
21.2037 | -106.773
21.2074 | -106.7655
21.2111 | -106.7572
21.2148 | -106.7471
21.2185 | -106.7385
21.2222 | -106.7324
21.2259 | -106.727
21.2296 | -106.7227
21.2333 | -106.7188
21.237 | -106.7172
21.2407 | -106.715
21.2444 | -106.7137
21.2481 | -106.715
21.2519 | -106.7159
21.2556 | -106.7163
21.2593 | -106.7173
21.263 | -106.7182
21.2667 | -106.7206
21.2704 | -106.7241
21.2741 | -106.7284
21.2778 | -106.7305
21.2815 | -106.7335
21.2852 | -106.7373
21.2889 | -106.7408
21.2926 | -106.7448
21.2963 | -106.7488
21.3 | -106.7524
21.3037 | -106.7546
21.3074 | -106.7558
21.3111 | -106.7564
21.3148 | -106.7559
21.3185 | -106.7541
21.3222 | -106.7544
21.3259 | -106.7568
21.3296 | -106.756
21.3333 | -106.7542
21.337 | -106.7519
21.3407 | -106.747
21.3444 | -106.7438
21.3481 | -106.7412
21.3519 | -106.7378
21.3556 | -106.7344
21.3593 | -106.7299
21.363 | -106.7228
21.3667 | -106.7159
21.3704 | -106.7088
21.3741 | -106.7032
21.3778 | -106.6989
21.3815 | -106.6962
21.3852 | -106.6925
21.3889 | -106.6876
21.3926 | -106.6842
21.3963 | -106.6802
21.4 | -106.6769
21.4037 | -106.6738
21.4074 | -106.6703
21.4111 | -106.6671
21.4148 | -106.6656
21.4185 | -106.6623
21.4222 | -106.6604
21.4259 | -106.6606
21.4296 | -106.6613
21.4333 | -106.661
21.437 | -106.6612
21.4407 | -106.6621
21.4444 | -106.6631
21.4481 | -106.6651
21.4519 | -106.6677
21.4556 | -106.6729
21.4593 | -106.6782
21.463 | -106.6836
21.4667 | -106.6889
21.4704 | -106.6937
21.4741 | -106.6995
21.4778 | -106.7078
21.4815 | -106.7155
21.4852 | -106.7261
21.4889 | -106.7364
21.4926 | -106.7468
21.4963 | -106.7577
21.5 | -106.7689
21.5037 | -106.7798
21.5074 | -106.7909
21.5111 | -106.8026
21.5148 | -106.8165
21.5185 | -106.8313
21.5222 | -106.8461
21.5259 | -106.8601
21.5296 | -106.8753
21.5333 | -106.8866
21.537 | -106.8987
21.5407 | -106.9109
21.5444 | -106.9233
21.5481 | -106.9328
21.5519 | -106.9405
21.5556 | -106.9516
21.5593 | -106.9639
21.563 | -106.9728
21.5667 | -106.981
21.5704 | -106.9879
21.5741 | -106.9954
21.5778 | -107.0037
21.5815 | -107.0113
21.5852 | -107.0148
21.5889 | -107.0171
21.5926 | -107.021
21.5963 | -107.0227
21.6 | -107.0239
21.6037 | -107.0246
21.6074 | -107.0251
21.6111 | -107.0247
21.6148 | -107.024
21.6185 | -107.0225
21.6222 | -107.0222
21.6259 | -107.0205
21.6296 | -107.0171
21.6333 | -107.015
21.637 | -107.0102
21.6407 | -107.0024
21.6444 | -106.9976
21.6481 | -106.9921
21.6519 | -106.9862
21.6556 | -106.9797
21.6593 | -106.9738
21.663 | -106.9692
21.6667 | -106.9642
21.6704 | -106.9582
21.6741 | -106.9523
21.6778 | -106.9464
21.6815 | -106.9394
21.6852 | -106.9329
21.6889 | -106.9272
21.6926 | -106.9218
21.6963 | -106.9181
21.7 | -106.9151
21.7037 | -106.9108
21.7074 | -106.9065
21.7111 | -106.9032
21.7148 | -106.8994
21.7185 | -106.8944
21.7222 | -106.8909
21.7259 | -106.8894
21.7296 | -106.888
21.7333 | -106.8849
21.737 | -106.8818
21.7407 | -106.8821
21.7444 | -106.8813
21.7481 | -106.8786
21.7519 | -106.8759
21.7556 | -106.8752
21.7593 | -106.8717
21.763 | -106.8671
21.7667 | -106.8654
21.7704 | -106.8631
21.7741 | -106.8611
21.7778 | -106.8596
21.7815 | -106.8572
21.7852 | -106.853
21.7889 | -106.8497
21.7926 | -106.8473
21.7963 | -106.846
21.8 | -106.8426
21.8037 | -106.8388
21.8074 | -106.8345
21.8111 | -106.8324
21.8148 | -106.8287
21.8185 | -106.8242
21.8222 | -106.8188
21.8259 | -106.8159
21.8296 | -106.8113
21.8333 | -106.8039
21.837 | -106.7963
21.8407 | -106.7938
21.8444 | -106.7893
21.8481 | -106.7829
21.8519 | -106.7782
21.8556 | -106.7743
21.8593 | -106.7706
21.863 | -106.7642
21.8667 | -106.7565
21.8704 | -106.749
21.8741 | -106.7429
21.8778 | -106.7361
21.8815 | -106.7283
21.8852 | -106.7223
21.8889 | -106.716
21.8926 | -106.7103
21.8963 | -106.7074
21.9 | -106.7024
21.9037 | -106.6967
21.9074 | -106.6921
21.9111 | -106.6854
21.9148 | -106.6788
21.9185 | -106.6707
21.9222 | -106.664
21.9259 | -106.6574
21.9296 | -106.6516
21.9333 | -106.6441
21.937 | -106.6369
21.9407 | -106.6295
21.9444 | -106.6219
21.9481 | -106.6154
21.9519 | -106.6109
21.9556 | -106.6062
21.9593 | -106.6006
21.963 | -106.5933
21.9667 | -106.5872
21.9704 | -106.5799
21.9741 | -106.5723
21.9778 | -106.5661
21.9815 | -106.5591
21.9852 | -106.554
21.9889 | -106.549
21.9926 | -106.5436
21.9963 | -106.5372
22 | -106.5322
22.0037 | -106.5277
22.0074 | -106.5216
22.0111 | -106.5151
22.0148 | -106.5109
22.0185 | -106.5053
22.0222 | -106.5014
22.0259 | -106.4985
22.0296 | -106.4947
22.0333 | -106.4891
22.037 | -106.4834
22.0407 | -106.4807
22.0444 | -106.4799
22.0481 | -106.4791
22.0519 | -106.4747
22.0556 | -106.4699
22.0593 | -106.4683
22.063 | -106.4672
22.0667 | -106.4643
22.0704 | -106.4626
22.0741 | -106.4627
22.0778 | -106.4599
22.0815 | -106.4569
22.0852 | -106.4551
22.0889 | -106.4534
22.0926 | -106.4513
22.0963 | -106.4494
22.1 | -106.4483
22.1037 | -106.4456
22.1074 | -106.4427
22.1111 | -106.442
22.1148 | -106.4415
22.1185 | -106.4382
22.1222 | -106.4338
22.1259 | -106.4316
22.1296 | -106.4303
22.1333 | -106.4287
22.137 | -106.4265
22.1407 | -106.4251
22.1444 | -106.4247
22.1481 | -106.424
22.1519 | -106.4214
22.1556 | -106.4178
22.1593 | -106.4152
22.163 | -106.4118
22.1667 | -106.4086
22.1704 | -106.4057
22.1741 | -106.405
22.1778 | -106.4051
22.1815 | -106.406
22.1852 | -106.4038
22.1889 | -106.4031
22.1926 | -106.4034
22.1963 | -106.4023
22.2 | -106.4024
22.2037 | -106.4034
22.2074 | -106.4032
22.2111 | -106.4046
22.2148 | -106.4068
22.2185 | -106.4075
22.2222 | -106.4099
22.2259 | -106.4136
22.2296 | -106.4157
22.2333 | -106.4178
22.237 | -106.421
22.2407 | -106.4256
22.2444 | -106.4303
22.2481 | -106.4371
22.2519 | -106.4435
22.2556 | -106.449
22.2593 | -106.4527
22.263 | -106.4587
22.2667 | -106.4644
22.2704 | -106.4689
22.2741 | -106.4747
22.2778 | -106.4814
22.2815 | -106.4867
22.2852 | -106.4921
22.2889 | -106.4974
22.2926 | -106.5035
22.2963 | -106.5102
22.3 | -106.5134
22.3037 | -106.5173
22.3074 | -106.522
22.3111 | -106.5255
22.3148 | -106.5282
22.3185 | -106.5313
22.3222 | -106.5342
22.3259 | -106.5374
22.3296 | -106.5398
22.3333 | -106.5425
22.337 | -106.5422
22.3407 | -106.5394
22.3444 | -106.5373
22.3481 | -106.5348
22.3519 | -106.5285
22.3556 | -106.5167
22.3593 | -106.505
22.363 | -106.4958
22.3667 | -106.4895
22.3704 | -106.4816
22.3741 | -106.4709
22.3778 | -106.4612
22.3815 | -106.45
22.3852 | -106.4381
22.3889 | -106.4252
22.3926 | -106.4128
22.3963 | -106.3995
22.4 | -106.3845
22.4037 | -106.3688
22.4074 | -106.3522
22.4111 | -106.3351
22.4148 | -106.3172
22.4185 | -106.2999
22.4222 | -106.2802
22.4259 | -106.2608
22.4296 | -106.2412
22.4333 | -106.2242
22.437 | -106.2077
22.4407 | -106.1902
22.4444 | -106.1716
22.4481 | -106.1527
22.4519 | -106.1344
22.4556 | -106.1179
22.4593 | -106.1031
22.463 | -106.0877
22.4667 | -106.0719
22.4704 | -106.058
22.4741 | -106.0482
22.4778 | -106.0393
22.4815 | -106.0304
22.4852 | -106.0222
22.4889 | -106.0164
22.4926 | -106.0126
22.4963 | -106.01
22.5 | -106.0093
22.5037 | -106.0124
22.5074 | -106.0144
22.5111 | -106.0178
22.5148 | -106.0247
22.5185 | -106.0329
22.5222 | -106.0437
22.5259 | -106.0563
22.5296 | -106.0705
22.5333 | -106.0862
22.537 | -106.1046
22.5407 | -106.1247
22.5444 | -106.1456
22.5481 | -106.17
22.5519 | -106.1953
22.5556 | -106.2201
22.5593 | -106.2459
22.563 | -106.2727
22.5667 | -106.2995
22.5704 | -106.3277
22.5741 | -106.3575
22.5778 | -106.3881
22.5815 | -106.4189
22.5852 | -106.4506
22.5889 | -106.4825
22.5926 | -106.5118
22.5963 | -106.5408
22.6 | -106.5706
22.6037 | -106.6008
22.6074 | -106.6305
22.6111 | -106.6574
22.6148 | -106.6828
22.6185 | -106.7104
22.6222 | -106.7347
22.6259 | -106.7574
22.6296 | -106.781
22.6333 | -106.8007
22.637 | -106.8198
22.6407 | -106.8382
22.6444 | -106.854
22.6481 | -106.8678
22.6519 | -106.8794
22.6556 | -106.8912
22.6593 | -106.9039
22.663 | -106.9141
22.6667 | -106.9221
22.6704 | -106.9311
22.6741 | -106.941
22.6778 | -106.9473
22.6815 | -106.9477
22.6852 | -106.9441
22.6889 | -106.9427
22.6926 | -106.9419
22.6963 | -106.9363
22.7 | -106.9318
22.7037 | -106.9277
22.7074 | -106.9206
22.7111 | -106.9138
22.7148 | -106.9065
22.7185 | -106.8982
22.7222 | -106.8903
22.7259 | -106.883
22.7296 | -106.8755
22.7333 | -106.8659
22.737 | -106.855
22.7407 | -106.8447
22.7444 | -106.8344
22.7481 | -106.8253
22.7519 | -106.8152
22.7556 | -106.804
22.7593 | -106.7947
22.763 | -106.7873
22.7667 | -106.7796
22.7704 | -106.7706
22.7741 | -106.761
22.7778 | -106.7501
22.7815 | -106.7398
22.7852 | -106.7314
22.7889 | -106.7231
22.7926 | -106.7128
22.7963 | -106.7017
22.8 | -106.6923
22.8037 | -106.6822
22.8074 | -106.6728
22.8111 | -106.6654
22.8148 | -106.6553
22.8185 | -106.6446
22.8222 | -106.6355
22.8259 | -106.6266
22.8296 | -106.6187
22.8333 | -106.6109
22.837 | -106.6015
22.8407 | -106.5939
22.8444 | -106.587
22.8481 | -106.5791
22.8519 | -106.5711
22.8556 | -106.5637
22.8593 | -106.5585
22.863 | -106.5521
22.8667 | -106.5452
22.8704 | -106.5382
22.8741 | -106.5334
22.8778 | -106.5279
22.8815 | -106.5226
22.8852 | -106.5167
22.8889 | -106.5121
22.8926 | -106.5092
22.8963 | -106.5066
22.9 | -106.5047
22.9037 | -106.5024
22.9074 | -106.5016
22.9111 | -106.5014
22.9148 | -106.5036
22.9185 | -106.5064
22.9222 | -106.5092
22.9259 | -106.5125
22.9296 | -106.5158
22.9333 | -106.52
22.937 | -106.5256
22.9407 | -106.5302
22.9444 | -106.5364
22.9481 | -106.5434
22.9519 | -106.5504
22.9556 | -106.5603
22.9593 | -106.5684
22.963 | -106.5775
22.9667 | -106.5846
22.9704 | -106.592
22.9741 | -106.6035
22.9778 | -106.6159
22.9815 | -106.6265
22.9852 | -106.6352
22.9889 | -106.6458
22.9926 | -106.6565
22.9963 | -106.6662
23 | -106.6762
23.0037 | -106.6863
23.0074 | -106.6977
23.0111 | -106.7084
23.0148 | -106.7157
23.0185 | -106.7231
23.0222 | -106.7309
23.0259 | -106.7386
23.0296 | -106.7443
23.0333 | -106.7517
23.037 | -106.7598
23.0407 | -106.7657
23.0444 | -106.7688
23.0481 | -106.7716
23.0519 | -106.7733
23.0556 | -106.7755
23.0593 | -106.7775
23.063 | -106.7807
23.0667 | -106.783
23.0704 | -106.7818
23.0741 | -106.7808
23.0778 | -106.7798
23.0815 | -106.7783
23.0852 | -106.7762
23.0889 | -106.7728
23.0926 | -106.7683
23.0963 | -106.7657
23.1 | -106.764
23.1037 | -106.7601
23.1074 | -106.7544
23.1111 | -106.7503
23.1148 | -106.7477
23.1185 | -106.7461
23.1222 | -106.7408
23.1259 | -106.7305
23.1296 | -106.7241
23.1333 | -106.7193
23.137 | -106.7124
23.1407 | -106.7064
23.1444 | -106.6998
23.1481 | -106.6937
23.1519 | -106.6871
23.1556 | -106.6805
23.1593 | -106.6737
23.163 | -106.6672
23.1667 | -106.6588
23.1704 | -106.6493
23.1741 | -106.6428
23.1778 | -106.6361
23.1815 | -106.6279
23.1852 | -106.6175
23.1889 | -106.6105
23.1926 | -106.6062
23.1963 | -106.6001
23.2 | -106.5921
23.2037 | -106.5849
23.2074 | -106.579
23.2111 | -106.5742
23.2148 | -106.5683
23.2185 | -106.5621
23.2222 | -106.5567
23.2259 | -106.553
23.2296 | -106.5471
23.2333 | -106.5401
23.237 | -106.5352
23.2407 | -106.5301
23.2444 | -106.5256
23.2481 | -106.522
23.2519 | -106.5186
23.2556 | -106.5143
23.2593 | -106.5108
23.263 | -106.5077
23.2667 | -106.5059
23.2704 | -106.5052
23.2741 | -106.5049
23.2778 | -106.5037
23.2815 | -106.5029
23.2852 | -106.503
23.2889 | -106.502
23.2926 | -106.5025
23.2963 | -106.5041
23.3 | -106.504
23.3037 | -106.5036
23.3074 | -106.5036
23.3111 | -106.5049
23.3148 | -106.5071
23.3185 | -106.5093
23.3222 | -106.5135
23.3259 | -106.5201
23.3296 | -106.5255
23.3333 | -106.5308
23.337 | -106.5367
23.3407 | -106.5429
23.3444 | -106.5497
23.3481 | -106.5557
23.3519 | -106.5634
23.3556 | -106.5705
23.3593 | -106.5797
23.363 | -106.5886
23.3667 | -106.5963
23.3704 | -106.6042
23.3741 | -106.6137
23.3778 | -106.6247
23.3815 | -106.6348
23.3852 | -106.6432
23.3889 | -106.6535
23.3926 | -106.6628
23.3963 | -106.6718
23.4 | -106.6817
23.4037 | -106.6912
23.4074 | -106.7001
23.4111 | -106.7074
23.4148 | -106.713
23.4185 | -106.7187
23.4222 | -106.7267
23.4259 | -106.7325
23.4296 | -106.7342
23.4333 | -106.7371
23.437 | -106.7423
23.4407 | -106.7434
23.4444 | -106.7442
23.4481 | -106.7447
23.4519 | -106.7479
23.4556 | -106.7515
23.4593 | -106.7528
23.463 | -106.7513
23.4667 | -106.7488
23.4704 | -106.7475
23.4741 | -106.7485
23.4778 | -106.7462
23.4815 | -106.7451
23.4852 | -106.7435
23.4889 | -106.7385
23.4926 | -106.7331
23.4963 | -106.7292
23.5 | -106.7245
23.5037 | -106.7181
23.5074 | -106.7146
23.5111 | -106.7093
23.5148 | -106.7058
23.5185 | -106.7026
23.5222 | -106.6983
23.5259 | -106.692
23.5296 | -106.6863
23.5333 | -106.6825
23.537 | -106.6786
23.5407 | -106.6739
23.5444 | -106.6684
23.5481 | -106.6624
23.5519 | -106.6569
23.5556 | -106.6524
23.5593 | -106.6475
23.563 | -106.6439
23.5667 | -106.6407
23.5704 | -106.6364
23.5741 | -106.6314
23.5778 | -106.6263
23.5815 | -106.6232
23.5852 | -106.62
23.5889 | -106.6162
23.5926 | -106.6105
23.5963 | -106.6053
23.6 | -106.6004
23.6037 | -106.595
23.6074 | -106.5905
23.6111 | -106.587
23.6148 | -106.5819
23.6185 | -106.5786
23.6222 | -106.5755
23.6259 | -106.5731
23.6296 | -106.5689
23.6333 | -106.5644
23.637 | -106.5587
23.6407 | -106.5534
23.6444 | -106.5492
23.6481 | -106.545
23.6519 | -106.5397
23.6556 | -106.5342
23.6593 | -106.527
23.663 | -106.522
23.6667 | -106.5182
23.6704 | -106.5153
23.6741 | -106.5139
23.6778 | -106.5125
23.6815 | -106.5105
23.6852 | -106.5089
23.6889 | -106.5068
23.6926 | -106.5035
23.6963 | -106.502
23.7 | -106.5012
23.7037 | -106.5015
23.7074 | -106.503
23.7111 | -106.5026
23.7148 | -106.5025
23.7185 | -106.5033
23.7222 | -106.5027
23.7259 | -106.5031
23.7296 | -106.5056
23.7333 | -106.5091
23.737 | -106.5117
23.7407 | -106.5148
23.7444 | -106.5192
23.7481 | -106.526
23.7519 | -106.5317
23.7556 | -106.5365
23.7593 | -106.5438
23.763 | -106.5533
23.7667 | -106.5619
23.7704 | -106.5692
23.7741 | -106.5789
23.7778 | -106.588
23.7815 | -106.5963
23.7852 | -106.6053
23.7889 | -106.6154
23.7926 | -106.6247
23.7963 | -106.6362
23.8 | -106.6461
23.8037 | -106.6542
23.8074 | -106.6623
23.8111 | -106.6727
23.8148 | -106.6815
23.8185 | -106.6918
23.8222 | -106.7046
23.8259 | -106.7163
23.8296 | -106.7277
23.8333 | -106.741
23.837 | -106.7529
23.8407 | -106.7625
23.8444 | -106.7727
23.8481 | -106.7853
23.8519 | -106.7968
23.8556 | -106.8073
23.8593 | -106.8173
23.863 | -106.8289
23.8667 | -106.8427
23.8704 | -106.8537
23.8741 | -106.8634
23.8778 | -106.874
23.8815 | -106.8857
23.8852 | -106.8974
23.8889 | -106.907
23.8926 | -106.9159
23.8963 | -106.9243
23.9 | -106.9321
23.9037 | -106.9393
23.9074 | -106.9468
23.9111 | -106.9527
23.9148 | -106.9585
23.9185 | -106.9638
23.9222 | -106.9696
23.9259 | -106.977
23.9296 | -106.9852
23.9333 | -106.9923
23.937 | -106.9965
23.9407 | -106.9989
23.9444 | -107.0008
23.9481 | -107.002
23.9519 | -107.0025
23.9556 | -107.0033
23.9593 | -107.0055
23.963 | -107.0078
23.9667 | -107.0083
23.9704 | -107.0088
23.9741 | -107.0107
23.9778 | -107.01
23.9815 | -107.0064
23.9852 | -107.0043
23.9889 | -107.0032
23.9926 | -107.0003
23.9963 | -106.9991
24 | -106.9976
24.0037 | -106.9954
24.0074 | -106.9906
24.0111 | -106.9857
24.0148 | -106.982
24.0185 | -106.9792
24.0222 | -106.9746
24.0259 | -106.9706
24.0296 | -106.9665
24.0333 | -106.9619
24.037 | -106.9577
24.0407 | -106.954
24.0444 | -106.9461
24.0481 | -106.9398
24.0519 | -106.9325
24.0556 | -106.9284
24.0593 | -106.9253
24.063 | -106.921
24.0667 | -106.9141
24.0704 | -106.9072
24.0741 | -106.9009
24.0778 | -106.8951
24.0815 | -106.8885
24.0852 | -106.881
24.0889 | -106.8739
24.0926 | -106.8679
24.0963 | -106.8613
24.1 | -106.8538
24.1037 | -106.8464
24.1074 | -106.8393
24.1111 | -106.8327
24.1148 | -106.8259
24.1185 | -106.8184
24.1222 | -106.813
24.1259 | -106.8062
24.1296 | -106.8006
24.1333 | -106.7962
24.137 | -106.7893
24.1407 | -106.783
24.1444 | -106.7788
24.1481 | -106.774
24.1519 | -106.7698
24.1556 | -106.767
24.1593 | -106.7641
24.163 | -106.7608
24.1667 | -106.7574
24.1704 | -106.7533
24.1741 | -106.7489
24.1778 | -106.747
24.1815 | -106.7449
24.1852 | -106.7422
24.1889 | -106.7387
24.1926 | -106.7367
24.1963 | -106.7364
24.2 | -106.7359
24.2037 | -106.735
24.2074 | -106.7343
24.2111 | -106.7355
24.2148 | -106.7373
24.2185 | -106.7374
24.2222 | -106.7367
24.2259 | -106.739
24.2296 | -106.7434
24.2333 | -106.7472
24.237 | -106.7505
24.2407 | -106.7533
24.2444 | -106.7556
24.2481 | -106.7582
24.2519 | -106.7625
24.2556 | -106.7654
24.2593 | -106.7634
24.263 | -106.7571
24.2667 | -106.7491
24.2704 | -106.7425
24.2741 | -106.7378
24.2778 | -106.7359
24.2815 | -106.7342
24.2852 | -106.7313
24.2889 | -106.7281
24.2926 | -106.7258
24.2963 | -106.7214
24.3 | -106.7159
24.3037 | -106.709
24.3074 | -106.7025
24.3111 | -106.6952
24.3148 | -106.6859
24.3185 | -106.6761
24.3222 | -106.6677
24.3259 | -106.6568
24.3296 | -106.6446
24.3333 | -106.6313
24.337 | -106.618
24.3407 | -106.6036
24.3444 | -106.5886
24.3481 | -106.5745
24.3519 | -106.5607
24.3556 | -106.5468
24.3593 | -106.5325
24.363 | -106.5188
24.3667 | -106.5037
24.3704 | -106.4872
24.3741 | -106.471
24.3778 | -106.4544
24.3815 | -106.439
24.3852 | -106.4254
24.3889 | -106.4117
24.3926 | -106.3989
24.3963 | -106.3867
24.4 | -106.3737
24.4037 | -106.3615
24.4074 | -106.3497
24.4111 | -106.3393
24.4148 | -106.3305
24.4185 | -106.323
24.4222 | -106.317
24.4259 | -106.3114
24.4296 | -106.3077
24.4333 | -106.3016
24.437 | -106.2984
24.4407 | -106.2989
24.4444 | -106.2997
24.4481 | -106.2999
24.4519 | -106.3014
24.4556 | -106.3053
24.4593 | -106.3093
24.463 | -106.3141
24.4667 | -106.3187
24.4704 | -106.3247
24.4741 | -106.3314
24.4778 | -106.3386
24.4815 | -106.3478
24.4852 | -106.3601
24.4889 | -106.372
24.4926 | -106.3841
24.4963 | -106.3944
24.5 | -106.402
24.5037 | -106.4118
24.5074 | -106.4221
24.5111 | -106.4331
24.5148 | -106.4444
24.5185 | -106.4556
24.5222 | -106.4659
24.5259 | -106.4769
24.5296 | -106.4878
24.5333 | -106.4994
24.537 | -106.5094
24.5407 | -106.5178
24.5444 | -106.5261
24.5481 | -106.5327
24.5519 | -106.5402
24.5556 | -106.5467
24.5593 | -106.5533
24.563 | -106.5602
24.5667 | -106.5667
24.5704 | -106.5656
24.5741 | -106.5625
24.5778 | -106.5334
24.5815 | -106.5106
24.5852 | -106.4653
24.5889 | -106.3919
24.5926 | -106.297
24.5963 | -106.2212
24.6 | -106.1717
24.6037 | -106.1237
24.6074 | -106.0745
24.6111 | -106.0215
24.6148 | -105.9588
24.6185 | -105.8877
24.6222 | -105.8117
24.6259 | -105.7352
24.6296 | -105.6561
24.6333 | -105.573
24.637 | -105.4875
24.6407 | -105.3993
24.6444 | -105.3086
24.6481 | -105.2144
24.6519 | -105.1187
24.6556 | -105.0218
24.6593 | -104.9237
24.663 | -104.8242
24.6667 | -104.7236
24.6704 | -104.6213
24.6741 | -104.519
24.6778 | -104.4176
24.6815 | -104.3174
24.6852 | -104.2168
24.6889 | -104.1176
24.6926 | -104.0196
24.6963 | -103.924
24.7 | -103.8305
24.7037 | -103.7406
24.7074 | -103.6541
24.7111 | -103.5711
24.7148 | -103.4917
24.7185 | -103.416
24.7222 | -103.3445
24.7259 | -103.2784
24.7296 | -103.2175
24.7333 | -103.1615
24.737 | -103.1107
24.7407 | -103.0646
24.7444 | -103.0234
24.7481 | -102.9874
24.7519 | -102.9574
24.7556 | -102.9346
24.7593 | -102.9175
24.763 | -102.9053
24.7667 | -102.8984
24.7704 | -102.8972
24.7741 | -102.9024
24.7778 | -102.9137
24.7815 | -102.932
24.7852 | -102.9544
24.7889 | -102.9814
24.7926 | -103.0147
24.7963 | -103.0544
24.8 | -103.0994
24.8037 | -103.148
24.8074 | -103.202
24.8111 | -103.261
24.8148 | -103.3245
24.8185 | -103.392
24.8222 | -103.4631
24.8259 | -103.5371
24.8296 | -103.6143
24.8333 | -103.6944
24.837 | -103.7773
24.8407 | -103.8627
24.8444 | -103.9506
24.8481 | -104.0404
24.8519 | -104.1305
24.8556 | -104.2217
24.8593 | -104.3139
24.863 | -104.4061
24.8667 | -104.4964
24.8704 | -104.5863
24.8741 | -104.6748
24.8778 | -104.7622
24.8815 | -104.8471
24.8852 | -104.9312
24.8889 | -105.0135
24.8926 | -105.0931
24.8963 | -105.1699
24.9 | -105.2427
24.9037 | -105.3135
24.9074 | -105.3816
24.9111 | -105.4468
24.9148 | -105.5096
24.9185 | -105.573
24.9222 | -105.6327
24.9259 | -105.6927
24.9296 | -105.751
24.9333 | -105.8097
24.937 | -105.8793
24.9407 | -105.9636
24.9444 | -106.0522
24.9481 | -106.1076
24.9519 | -106.1297
24.9556 | -106.1458
24.9593 | -106.1594
24.963 | -106.1749
24.9667 | -106.1947
24.9704 | -106.215
24.9741 | -106.2314
24.9778 | -106.2409
24.9815 | -106.2514
24.9852 | -106.2622
24.9889 | -106.2707
24.9926 | -106.2782
24.9963 | -106.2862
25 | -106.2953
25.0037 | -106.304
25.0074 | -106.3131
25.0111 | -106.3223
25.0148 | -106.3322
25.0185 | -106.3423
25.0222 | -106.3525
25.0259 | -106.3619
25.0296 | -106.37
25.0333 | -106.3783
25.037 | -106.3892
25.0407 | -106.3991
25.0444 | -106.4089
25.0481 | -106.4171
25.0519 | -106.4268
25.0556 | -106.4354
25.0593 | -106.442
25.063 | -106.4481
25.0667 | -106.456
25.0704 | -106.4654
25.0741 | -106.4725
25.0778 | -106.4774
25.0815 | -106.4816
25.0852 | -106.4846
25.0889 | -106.488
25.0926 | -106.4936
25.0963 | -106.4991
25.1 | -106.5024
25.1037 | -106.5037
25.1074 | -106.5047
25.1111 | -106.5033
25.1148 | -106.5038
25.1185 | -106.5045
25.1222 | -106.5042
25.1259 | -106.5053
25.1296 | -106.5067
25.1333 | -106.506
25.137 | -106.5029
25.1407 | -106.5019
25.1444 | -106.4992
25.1481 | -106.4952
25.1519 | -106.4917
25.1556 | -106.4877
25.1593 | -106.4833
25.163 | -106.4796
25.1667 | -106.4768
25.1704 | -106.4717
25.1741 | -106.4671
25.1778 | -106.4638
25.1815 | -106.4616
25.1852 | -106.4584
25.1889 | -106.4543
25.1926 | -106.4498
25.1963 | -106.4457
25.2 | -106.4421
25.2037 | -106.4397
25.2074 | -106.4367
25.2111 | -106.4341
25.2148 | -106.4305
25.2185 | -106.4253
25.2222 | -106.4219
25.2259 | -106.4189
25.2296 | -106.4181
25.2333 | -106.4169
25.237 | -106.4154
25.2407 | -106.412
25.2444 | -106.4086
25.2481 | -106.406
25.2519 | -106.4048
25.2556 | -106.4033
25.2593 | -106.4012
25.263 | -106.3975
25.2667 | -106.3939
25.2704 | -106.392
25.2741 | -106.3888
25.2778 | -106.3856
25.2815 | -106.3823
25.2852 | -106.3796
25.2889 | -106.3778
25.2926 | -106.3771
25.2963 | -106.374
25.3 | -106.3707
25.3037 | -106.3683
25.3074 | -106.3649
25.3111 | -106.3616
25.3148 | -106.3579
25.3185 | -106.3515
25.3222 | -106.3467
25.3259 | -106.3412
25.3296 | -106.3353
25.3333 | -106.3297
25.337 | -106.3229
25.3407 | -106.3161
25.3444 | -106.31
25.3481 | -106.304
25.3519 | -106.298
25.3556 | -106.2921
25.3593 | -106.2862
25.363 | -106.2781
25.3667 | -106.2688
25.3704 | -106.2605
25.3741 | -106.2535
25.3778 | -106.2476
25.3815 | -106.2419
25.3852 | -106.2366
25.3889 | -106.2306
25.3926 | -106.2224
25.3963 | -106.2158
25.4 | -106.209
25.4037 | -106.2012
25.4074 | -106.192
25.4111 | -106.184
25.4148 | -106.1776
25.4185 | -106.1706
25.4222 | -106.1627
25.4259 | -106.1533
25.4296 | -106.145
25.4333 | -106.1375
25.437 | -106.1293
25.4407 | -106.1219
25.4444 | -106.116
25.4481 | -106.1079
25.4519 | -106.1009
25.4556 | -106.0957
25.4593 | -106.0906
25.463 | -106.0841
25.4667 | -106.077
25.4704 | -106.0705
25.4741 | -106.0641
25.4778 | -106.0577
25.4815 | -106.0518
25.4852 | -106.0461
25.4889 | -106.0403
25.4926 | -106.0355
25.4963 | -106.0318
25.5 | -106.0274
25.5037 | -106.0238
25.5074 | -106.0196
25.5111 | -106.0138
25.5148 | -106.0072
25.5185 | -106.0025
25.5222 | -105.997
25.5259 | -105.9908
25.5296 | -105.9867
25.5333 | -105.9823
25.537 | -105.977
25.5407 | -105.9728
25.5444 | -105.9681
25.5481 | -105.9619
25.5519 | -105.9563
25.5556 | -105.9502
25.5593 | -105.9463
25.563 | -105.9426
25.5667 | -105.9383
25.5704 | -105.9358
25.5741 | -105.9331
25.5778 | -105.9311
25.5815 | -105.9294
25.5852 | -105.927
25.5889 | -105.9245
25.5926 | -105.9231
25.5963 | -105.9219
25.6 | -105.9201
25.6037 | -105.9169
25.6074 | -105.9153
25.6111 | -105.9148
25.6148 | -105.9142
25.6185 | -105.9132
25.6222 | -105.9123
25.6259 | -105.9108
25.6296 | -105.908
25.6333 | -105.9074
25.637 | -105.9089
25.6407 | -105.9103
25.6444 | -105.9149
25.6481 | -105.9172
25.6519 | -105.9202
25.6556 | -105.9217
25.6593 | -105.9249
25.663 | -105.9281
25.6667 | -105.9313
25.6704 | -105.9342
25.6741 | -105.9366
25.6778 | -105.9402
25.6815 | -105.9468
25.6852 | -105.9524
25.6889 | -105.9569
25.6926 | -105.9627
25.6963 | -105.9677
25.7 | -105.9727
25.7037 | -105.9765
25.7074 | -105.9811
25.7111 | -105.9879
25.7148 | -105.9929
25.7185 | -105.9971
25.7222 | -106.0018
25.7259 | -106.0075
25.7296 | -106.0151
25.7333 | -106.0209
25.737 | -106.0263
25.7407 | -106.0329
25.7444 | -106.0392
25.7481 | -106.046
25.7519 | -106.0529
25.7556 | -106.0596
25.7593 | -106.0652
25.763 | -106.0716
25.7667 | -106.0805
25.7704 | -106.0885
25.7741 | -106.0928
25.7778 | -106.0979
25.7815 | -106.1043
25.7852 | -106.1094
25.7889 | -106.1146
25.7926 | -106.1186
25.7963 | -106.1225
25.8 | -106.1279
25.8037 | -106.1335
25.8074 | -106.1371
25.8111 | -106.1412
25.8148 | -106.144
25.8185 | -106.1472
25.8222 | -106.1512
25.8259 | -106.1534
25.8296 | -106.1562
25.8333 | -106.1581
25.837 | -106.1601
25.8407 | -106.1622
25.8444 | -106.1644
25.8481 | -106.1656
25.8519 | -106.1662
25.8556 | -106.1664
25.8593 | -106.1664
25.863 | -106.1658
25.8667 | -106.1656
25.8704 | -106.1676
25.8741 | -106.1697
25.8778 | -106.1709
25.8815 | -106.1721
25.8852 | -106.172
25.8889 | -106.1708
25.8926 | -106.1702
25.8963 | -106.1709
25.9 | -106.1702
25.9037 | -106.1693
25.9074 | -106.1703
25.9111 | -106.171
25.9148 | -106.172
25.9185 | -106.1722
25.9222 | -106.1716
25.9259 | -106.171
25.9296 | -106.1708
25.9333 | -106.1685
25.937 | -106.1683
25.9407 | -106.1699
25.9444 | -106.1711
25.9481 | -106.1722
25.9519 | -106.1732
25.9556 | -106.1744
25.9593 | -106.1773
25.963 | -106.1823
25.9667 | -106.1874
25.9704 | -106.1898
25.9741 | -106.1919
25.9778 | -106.195
25.9815 | -106.1987
25.9852 | -106.2039
25.9889 | -106.2098
25.9926 | -106.2166
25.9963 | -106.2213
26 | -106.2232
26.0037 | -106.2269
26.0074 | -106.2305
26.0111 | -106.2349
26.0148 | -106.2395
26.0185 | -106.2451
26.0222 | -106.2513
26.0259 | -106.2575
26.0296 | -106.2637
26.0333 | -106.2689
26.037 | -106.2742
26.0407 | -106.2809
26.0444 | -106.2871
26.0481 | -106.2917
26.0519 | -106.2952
26.0556 | -106.3001
26.0593 | -106.3049
26.063 | -106.3106
26.0667 | -106.3146
26.0704 | -106.3186
26.0741 | -106.323
26.0778 | -106.3246
26.0815 | -106.3277
26.0852 | -106.332
26.0889 | -106.335
26.0926 | -106.3371
26.0963 | -106.3372
26.1 | -106.3368
26.1037 | -106.3359
26.1074 | -106.3337
26.1111 | -106.3328
26.1148 | -106.3318
26.1185 | -106.3317
26.1222 | -106.3333
26.1259 | -106.3336
26.1296 | -106.3329
26.1333 | -106.3317
26.137 | -106.3325
26.1407 | -106.3355
26.1444 | -106.3362
26.1481 | -106.3362
26.1519 | -106.3364
26.1556 | -106.3374
26.1593 | -106.3394
26.163 | -106.3401
26.1667 | -106.3385
26.1704 | -106.3363
26.1741 | -106.3337
26.1778 | -106.3327
26.1815 | -106.3339
26.1852 | -106.3345
26.1889 | -106.3342
26.1926 | -106.3321
26.1963 | -106.3327
26.2 | -106.3331
26.2037 | -106.3344
26.2074 | -106.3359
26.2111 | -106.337
26.2148 | -106.3385
26.2185 | -106.3407
26.2222 | -106.3433
26.2259 | -106.3461
26.2296 | -106.3487
26.2333 | -106.3506
26.237 | -106.3529
26.2407 | -106.354
26.2444 | -106.355
26.2481 | -106.3578
26.2519 | -106.3591
26.2556 | -106.3624
26.2593 | -106.3671
26.263 | -106.3694
26.2667 | -106.3701
26.2704 | -106.373
26.2741 | -106.3756
26.2778 | -106.3777
26.2815 | -106.3813
26.2852 | -106.384
26.2889 | -106.3874
26.2926 | -106.3901
26.2963 | -106.3942
26.3 | -106.3992
26.3037 | -106.4029
26.3074 | -106.4063
26.3111 | -106.4096
26.3148 | -106.4139
26.3185 | -106.418
26.3222 | -106.4215
26.3259 | -106.4243
26.3296 | -106.4267
26.3333 | -106.4297
26.337 | -106.4336
26.3407 | -106.4363
26.3444 | -106.4401
26.3481 | -106.4447
26.3519 | -106.4485
26.3556 | -106.4526
26.3593 | -106.4559
26.363 | -106.4583
26.3667 | -106.4617
26.3704 | -106.4671
26.3741 | -106.4714
26.3778 | -106.475
26.3815 | -106.4787
26.3852 | -106.482
26.3889 | -106.4837
26.3926 | -106.4853
26.3963 | -106.4873
26.4 | -106.4894
26.4037 | -106.4918
26.4074 | -106.4958
26.4111 | -106.499
26.4148 | -106.5027
26.4185 | -106.507
26.4222 | -106.511
26.4259 | -106.514
26.4296 | -106.516
26.4333 | -106.5194
26.437 | -106.5237
26.4407 | -106.5276
26.4444 | -106.5295
26.4481 | -106.5322
26.4519 | -106.5346
26.4556 | -106.5358
26.4593 | -106.5384
26.463 | -106.5415
26.4667 | -106.5446
26.4704 | -106.5493
26.4741 | -106.5531
26.4778 | -106.5559
26.4815 | -106.5571
26.4852 | -106.5576
26.4889 | -106.5585
26.4926 | -106.5596
26.4963 | -106.561
26.5 | -106.5627
26.5037 | -106.5646
26.5074 | -106.5649
26.5111 | -106.5652
26.5148 | -106.5674
26.5185 | -106.5697
26.5222 | -106.5718
26.5259 | -106.5743
26.5296 | -106.5758
26.5333 | -106.5778
26.537 | -106.5804
26.5407 | -106.5822
26.5444 | -106.586
26.5481 | -106.5897
26.5519 | -106.5933
26.5556 | -106.5962
26.5593 | -106.598
26.563 | -106.5999
26.5667 | -106.6031
26.5704 | -106.6057
26.5741 | -106.6083
26.5778 | -106.6098
26.5815 | -106.6115
26.5852 | -106.6132
26.5889 | -106.6137
26.5926 | -106.6153
26.5963 | -106.6168
26.6 | -106.6181
26.6037 | -106.6199
26.6074 | -106.6242
26.6111 | -106.6281
26.6148 | -106.6309
26.6185 | -106.6321
26.6222 | -106.6339
26.6259 | -106.6383
26.6296 | -106.6411
26.6333 | -106.6445
26.637 | -106.649
26.6407 | -106.6533
26.6444 | -106.6569
26.6481 | -106.6588
26.6519 | -106.6621
26.6556 | -106.664
26.6593 | -106.6675
26.663 | -106.6716
26.6667 | -106.6744
26.6704 | -106.6752
26.6741 | -106.676
26.6778 | -106.6778
26.6815 | -106.6796
26.6852 | -106.6813
26.6889 | -106.6825
26.6926 | -106.6831
26.6963 | -106.6837
26.7 | -106.6858
26.7037 | -106.6886
26.7074 | -106.6905
26.7111 | -106.6911
26.7148 | -106.6916
26.7185 | -106.6948
26.7222 | -106.6944
26.7259 | -106.6922
26.7296 | -106.6906
26.7333 | -106.6893
26.737 | -106.6889
26.7407 | -106.6891
26.7444 | -106.6886
26.7481 | -106.6868
26.7519 | -106.685
26.7556 | -106.6841
26.7593 | -106.6835
26.763 | -106.6823
26.7667 | -106.6807
26.7704 | -106.6795
26.7741 | -106.6781
26.7778 | -106.6768
26.7815 | -106.674
26.7852 | -106.6691
26.7889 | -106.6659
26.7926 | -106.6627
26.7963 | -106.6582
26.8 | -106.655
26.8037 | -106.6513
26.8074 | -106.6478
26.8111 | -106.6439
26.8148 | -106.6398
26.8185 | -106.6357
26.8222 | -106.6336
26.8259 | -106.6296
26.8296 | -106.6262
26.8333 | -106.6234
26.837 | -106.6196
26.8407 | -106.6163
26.8444 | -106.6126
26.8481 | -106.6093
26.8519 | -106.6069
26.8556 | -106.6038
26.8593 | -106.6006
26.863 | -106.5984
26.8667 | -106.5949
26.8704 | -106.592
26.8741 | -106.5902
26.8778 | -106.5894
26.8815 | -106.587
26.8852 | -106.5842
26.8889 | -106.5829
26.8926 | -106.5813
26.8963 | -106.5786
26.9 | -106.578
26.9037 | -106.5771
26.9074 | -106.5758
26.9111 | -106.5745
26.9148 | -106.5728
26.9185 | -106.5723
26.9222 | -106.5711
26.9259 | -106.5706
26.9296 | -106.5701
26.9333 | -106.569
26.937 | -106.5689
26.9407 | -106.5699
26.9444 | -106.5715
26.9481 | -106.5716
26.9519 | -106.5728
26.9556 | -106.5749
26.9593 | -106.5774
26.963 | -106.579
26.9667 | -106.5825
26.9704 | -106.5836
26.9741 | -106.585
26.9778 | -106.5877
26.9815 | -106.5903
26.9852 | -106.5924
26.9889 | -106.595
26.9926 | -106.598
26.9963 | -106.6017
27 | -106.6038
27.0037 | -106.6041
27.0074 | -106.605
27.0111 | -106.6051
27.0148 | -106.6045
27.0185 | -106.6053
27.0222 | -106.6063
27.0259 | -106.6069
27.0296 | -106.6068
27.0333 | -106.6088
27.037 | -106.6091
27.0407 | -106.6091
27.0444 | -106.6081
27.0481 | -106.6078
27.0519 | -106.6062
27.0556 | -106.6052
27.0593 | -106.6039
27.063 | -106.602
27.0667 | -106.5988
27.0704 | -106.5951
27.0741 | -106.5907
27.0778 | -106.5853
27.0815 | -106.5805
27.0852 | -106.5753
27.0889 | -106.5712
27.0926 | -106.5646
27.0963 | -106.558
27.1 | -106.5523
27.1037 | -106.5472
27.1074 | -106.5425
27.1111 | -106.538
27.1148 | -106.5329
27.1185 | -106.527
27.1222 | -106.5206
27.1259 | -106.5146
27.1296 | -106.5087
27.1333 | -106.5033
27.137 | -106.4953
27.1407 | -106.4882
27.1444 | -106.483
27.1481 | -106.4794
27.1519 | -106.4751
27.1556 | -106.4703
27.1593 | -106.4668
27.163 | -106.465
27.1667 | -106.4626
27.1704 | -106.4591
27.1741 | -106.4564
27.1778 | -106.4555
27.1815 | -106.4544
27.1852 | -106.4535
27.1889 | -106.4527
27.1926 | -106.4522
27.1963 | -106.4525
27.2 | -106.4535
27.2037 | -106.4532
27.2074 | -106.4541
27.2111 | -106.4567
27.2148 | -106.4593
27.2185 | -106.4617
27.2222 | -106.4646
27.2259 | -106.469
27.2296 | -106.4726
27.2333 | -106.4761
27.237 | -106.4807
27.2407 | -106.4874
27.2444 | -106.4937
27.2481 | -106.5029
27.2519 | -106.5084
27.2556 | -106.5141
27.2593 | -106.5207
27.263 | -106.5275
27.2667 | -106.5351
27.2704 | -106.5419
27.2741 | -106.5487
27.2778 | -106.5558
27.2815 | -106.5651
27.2852 | -106.5741
27.2889 | -106.5826
27.2926 | -106.591
27.2963 | -106.5986
27.3 | -106.6068
27.3037 | -106.6144
27.3074 | -106.6217
27.3111 | -106.6292
27.3148 | -106.6366
27.3185 | -106.6443
27.3222 | -106.6532
27.3259 | -106.6601
27.3296 | -106.6644
27.3333 | -106.6713
27.337 | -106.6776
27.3407 | -106.6823
27.3444 | -106.6857
27.3481 | -106.6904
27.3519 | -106.6935
27.3556 | -106.6974
27.3593 | -106.7014
27.363 | -106.7044
27.3667 | -106.7081
27.3704 | -106.7093
27.3741 | -106.7097
27.3778 | -106.71
27.3815 | -106.712
27.3852 | -106.7138
27.3889 | -106.7149
27.3926 | -106.7163
27.3963 | -106.7162
27.4 | -106.717
27.4037 | -106.717
27.4074 | -106.7145
27.4111 | -106.7136
27.4148 | -106.7117
27.4185 | -106.7088
27.4222 | -106.7076
27.4259 | -106.7057
27.4296 | -106.7043
27.4333 | -106.7027
27.437 | -106.7011
27.4407 | -106.6994
27.4444 | -106.6971
27.4481 | -106.694
27.4519 | -106.6915
27.4556 | -106.6908
27.4593 | -106.6892
27.463 | -106.6862
27.4667 | -106.6816
27.4704 | -106.6771
27.4741 | -106.6728
27.4778 | -106.6696
27.4815 | -106.6668
27.4852 | -106.662
27.4889 | -106.6588
27.4926 | -106.6558
27.4963 | -106.6513
27.5 | -106.6474
27.5037 | -106.6428
27.5074 | -106.6371
27.5111 | -106.6329
27.5148 | -106.6293
27.5185 | -106.6244
27.5222 | -106.6218
27.5259 | -106.6189
27.5296 | -106.6138
27.5333 | -106.6084
27.537 | -106.6061
27.5407 | -106.6039
27.5444 | -106.6002
27.5481 | -106.5957
27.5519 | -106.5931
27.5556 | -106.5895
27.5593 | -106.5858
27.563 | -106.5825
27.5667 | -106.5792
27.5704 | -106.5781
27.5741 | -106.5759
27.5778 | -106.5723
27.5815 | -106.569
27.5852 | -106.568
27.5889 | -106.5663
27.5926 | -106.564
27.5963 | -106.5584
27.6 | -106.5544
27.6037 | -106.5509
27.6074 | -106.5489
27.6111 | -106.5461
27.6148 | -106.5445
27.6185 | -106.5428
27.6222 | -106.5418
27.6259 | -106.5384
27.6296 | -106.5345
27.6333 | -106.5327
27.637 | -106.5317
27.6407 | -106.5296
27.6444 | -106.5277
27.6481 | -106.5257
27.6519 | -106.5246
27.6556 | -106.5248
27.6593 | -106.5255
27.663 | -106.525
27.6667 | -106.5252
27.6704 | -106.5242
27.6741 | -106.522
27.6778 | -106.5195
27.6815 | -106.5175
27.6852 | -106.5179
27.6889 | -106.5174
27.6926 | -106.5178
27.6963 | -106.5181
27.7 | -106.5173
27.7037 | -106.5165
27.7074 | -106.517
27.7111 | -106.5164
27.7148 | -106.5172
27.7185 | -106.5187
27.7222 | -106.5187
27.7259 | -106.5192
27.7296 | -106.5217
27.7333 | -106.524
27.737 | -106.5257
27.7407 | -106.5277
27.7444 | -106.5297
27.7481 | -106.5304
27.7519 | -106.5321
27.7556 | -106.5342
27.7593 | -106.5367
27.763 | -106.5403
27.7667 | -106.5449
27.7704 | -106.5495
27.7741 | -106.5538
27.7778 | -106.5581
27.7815 | -106.5631
27.7852 | -106.5663
27.7889 | -106.5723
27.7926 | -106.5772
27.7963 | -106.5808
27.8 | -106.586
27.8037 | -106.5898
27.8074 | -106.5943
27.8111 | -106.6
27.8148 | -106.6032
27.8185 | -106.6076
27.8222 | -106.6126
27.8259 | -106.617
27.8296 | -106.6219
27.8333 | -106.6256
27.837 | -106.6288
27.8407 | -106.6325
27.8444 | -106.6359
27.8481 | -106.6378
27.8519 | -106.6416
27.8556 | -106.6447
27.8593 | -106.6478
27.863 | -106.6508
27.8667 | -106.6519
27.8704 | -106.6526
27.8741 | -106.652
27.8778 | -106.6518
27.8815 | -106.6508
27.8852 | -106.65
27.8889 | -106.6505
27.8926 | -106.6499
27.8963 | -106.6498
27.9 | -106.6512
27.9037 | -106.65
27.9074 | -106.6487
27.9111 | -106.6483
27.9148 | -106.6457
27.9185 | -106.6415
27.9222 | -106.6367
27.9259 | -106.6263
27.9296 | -106.6119
27.9333 | -106.6013
27.937 | -106.5947
27.9407 | -106.5897
27.9444 | -106.5843
27.9481 | -106.5782
27.9519 | -106.5715
27.9556 | -106.5642
27.9593 | -106.5571
27.963 | -106.5483
27.9667 | -106.5406
27.9704 | -106.5343
27.9741 | -106.5279
27.9778 | -106.5203
27.9815 | -106.5132
27.9852 | -106.5059
27.9889 | -106.4995
27.9926 | -106.4914
27.9963 | -106.4825
28 | -106.4728
28.0037 | -106.4621