## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 288 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 294
Hood | 02-92B/P | 229
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1015.1 | hPa |  | PRC PRO 01
Humidity | 48.7 | %/Hr
Internal temperature of calibration bench | 23.57 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1019 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.4 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.02 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3121 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 14/09/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 19/09/2018
Technical Manager | Anthony HUE | Approval |  | 27/09/2018
CEO | Robert GLEMOT | Approval |  | 27/09/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4351 | 57.5951
0.012209 | -36.0067 | 38.3602
0.018313 | -34.9431 | 28.1287
0.024417 | -34.488 | 21.8959
0.030521 | -34.3652 | 18.1521
0.036626 | -34.2239 | 15.0825
0.04273 | -34.1324 | 12.7157
0.048834 | -34.0866 | 11.1802
0.054938 | -34.021 | 10.078
0.061043 | -33.9582 | 9.0617
0.067147 | -33.9423 | 8.2274
0.073251 | -33.9393 | 7.4576
0.079355 | -33.957 | 6.7847
0.08546 | -33.9746 | 6.1787
0.091564 | -33.9751 | 5.7145
0.097668 | -33.9617 | 5.4344
0.10377 | -33.9421 | 5.1826
0.10988 | -33.9263 | 4.9101
0.11598 | -33.9168 | 4.6337
0.12209 | -33.9176 | 4.3657
0.12819 | -33.9272 | 4.0906
0.13429 | -33.939 | 3.8043
0.1404 | -33.9437 | 3.527
0.1465 | -33.9463 | 3.2731
0.15261 | -33.9431 | 3.0583
0.15871 | -33.9355 | 2.9018
0.16482 | -33.926 | 2.8019
0.17092 | -33.9235 | 2.7527
0.17702 | -33.9215 | 2.6739
0.18313 | -33.9216 | 2.5896
0.18923 | -33.9225 | 2.4792
0.19534 | -33.9227 | 2.3581
0.20144 | -33.9228 | 2.2274
0.20754 | -33.9198 | 2.1177
0.21365 | -33.9188 | 2.0103
0.21975 | -33.9177 | 1.9151
0.22586 | -33.9177 | 1.8258
0.23196 | -33.9191 | 1.7469
0.23807 | -33.9221 | 1.6756
0.24417 | -33.926 | 1.61
0.25027 | -33.9296 | 1.5414
0.25638 | -33.9321 | 1.4751
0.26248 | -33.9342 | 1.4045
0.26859 | -33.9353 | 1.3369
0.27469 | -33.9359 | 1.2642
0.2808 | -33.9368 | 1.1886
0.2869 | -33.9377 | 1.1155
0.293 | -33.9385 | 1.0459
0.29911 | -33.9394 | 0.98296
0.30521 | -33.9403 | 0.92145
0.31132 | -33.9414 | 0.8799
0.31742 | -33.9427 | 0.83424
0.32353 | -33.9434 | 0.79595
0.32963 | -33.9446 | 0.76411
0.33573 | -33.9459 | 0.74151
0.34184 | -33.9475 | 0.7187
0.34794 | -33.9495 | 0.69603
0.35405 | -33.9515 | 0.67262
0.36015 | -33.9538 | 0.6488
0.36626 | -33.9556 | 0.62369
0.37236 | -33.9574 | 0.59042
0.37846 | -33.9592 | 0.54985
0.38457 | -33.9598 | 0.49758
0.39067 | -33.9607 | 0.45532
0.39678 | -33.9613 | 0.40667
0.40288 | -33.9616 | 0.3562
0.40899 | -33.9614 | 0.30677
0.41509 | -33.9609 | 0.25767
0.42119 | -33.9611 | 0.21303
0.4273 | -33.9613 | 0.16655
0.4334 | -33.9615 | 0.12016
0.43951 | -33.9621 | 0.075666
0.44561 | -33.9633 | 0.033883
0.45172 | -33.9652 | -0.00057007
0.45782 | -33.9664 | -0.039893
0.46392 | -33.9679 | -0.077709
0.47003 | -33.9693 | -0.11198
0.47613 | -33.9705 | -0.14697
0.48224 | -33.9712 | -0.18225
0.48834 | -33.9704 | -0.21531
0.49445 | -33.9703 | -0.25046
0.50055 | -33.9698 | -0.28161
0.50665 | -33.9692 | -0.31469
0.51276 | -33.9683 | -0.34678
0.51886 | -33.9672 | -0.37917
0.52497 | -33.9661 | -0.40287
0.53107 | -33.9655 | -0.43184
0.53717 | -33.9648 | -0.45959
0.54328 | -33.9641 | -0.48224
0.54938 | -33.9641 | -0.5049
0.55549 | -33.9643 | -0.52644
0.56159 | -33.9658 | -0.54006
0.5677 | -33.9664 | -0.5563
0.5738 | -33.9677 | -0.57455
0.5799 | -33.9688 | -0.59124
0.58601 | -33.9702 | -0.60343
0.59211 | -33.9721 | -0.61138
0.59822 | -33.9737 | -0.62217
0.60432 | -33.9754 | -0.63474
0.61043 | -33.9769 | -0.64398
0.61653 | -33.9782 | -0.65488
0.62263 | -33.9799 | -0.66474
0.62874 | -33.9804 | -0.67445
0.63484 | -33.9812 | -0.68387
0.64095 | -33.9817 | -0.69378
0.64705 | -33.9819 | -0.70327
0.65316 | -33.982 | -0.7147
0.65926 | -33.9819 | -0.72899
0.66536 | -33.9806 | -0.74545
0.67147 | -33.9802 | -0.75808
0.67757 | -33.9797 | -0.77178
0.68368 | -33.9788 | -0.79019
0.68978 | -33.9778 | -0.80696
0.69589 | -33.9769 | -0.82611
0.70199 | -33.976 | -0.85063
0.70809 | -33.9751 | -0.87046
0.7142 | -33.9744 | -0.8914
0.7203 | -33.9734 | -0.91222
0.72641 | -33.9728 | -0.93647
0.73251 | -33.9722 | -0.9585
0.73862 | -33.9727 | -0.98331
0.74472 | -33.9725 | -1.0057
0.75082 | -33.9725 | -1.0301
0.75693 | -33.9729 | -1.0533
0.76303 | -33.9735 | -1.077
0.76914 | -33.9752 | -1.0992
0.77524 | -33.9763 | -1.1204
0.78135 | -33.9776 | -1.1422
0.78745 | -33.979 | -1.1637
0.79355 | -33.9802 | -1.1849
0.79966 | -33.982 | -1.205
0.80576 | -33.9841 | -1.2231
0.81187 | -33.9857 | -1.244
0.81797 | -33.9874 | -1.2648
0.82408 | -33.9889 | -1.2844
0.83018 | -33.9905 | -1.3033
0.83628 | -33.9923 | -1.3207
0.84239 | -33.9933 | -1.3383
0.84849 | -33.9947 | -1.3585
0.8546 | -33.9959 | -1.3765
0.8607 | -33.997 | -1.3952
0.86681 | -33.998 | -1.412
0.87291 | -33.9988 | -1.4285
0.87901 | -33.999 | -1.444
0.88512 | -33.9996 | -1.4593
0.89122 | -34.0002 | -1.4757
0.89733 | -34.0004 | -1.4891
0.90343 | -34.0008 | -1.5044
0.90953 | -34.0009 | -1.5173
0.91564 | -34.0009 | -1.5317
0.92174 | -34.0012 | -1.5452
0.92785 | -34.0012 | -1.5577
0.93395 | -34.0013 | -1.5695
0.94006 | -34.0016 | -1.5811
0.94616 | -34.0016 | -1.5927
0.95226 | -34.0018 | -1.6045
0.95837 | -34.0021 | -1.6173
0.96447 | -34.0023 | -1.6292
0.97058 | -34.0024 | -1.6412
0.97668 | -34.0023 | -1.6537
0.98279 | -34.0023 | -1.6666
0.98889 | -34.0024 | -1.677
0.99499 | -34.0026 | -1.69
1.0011 | -34.0027 | -1.7003
1.0072 | -34.0026 | -1.7117
1.0133 | -34.0026 | -1.7222
1.0194 | -34.0025 | -1.7352
1.0255 | -34.0025 | -1.7492
1.0316 | -34.0023 | -1.7619
1.0377 | -34.0022 | -1.7749
1.0438 | -34.002 | -1.7874
1.0499 | -34.0015 | -1.8036
1.056 | -34.001 | -1.818
1.0621 | -34.0006 | -1.8326
1.0682 | -34.0004 | -1.8454
1.0743 | -34 | -1.8618
1.0805 | -33.9997 | -1.8764
1.0866 | -33.999 | -1.8923
1.0927 | -33.9984 | -1.9079
1.0988 | -33.998 | -1.9238
1.1049 | -33.9974 | -1.9391
1.111 | -33.9969 | -1.955
1.1171 | -33.9963 | -1.9704
1.1232 | -33.9959 | -1.9869
1.1293 | -33.9956 | -2.0022
1.1354 | -33.9955 | -2.0183
1.1415 | -33.9953 | -2.0352
1.1476 | -33.9949 | -2.0522
1.1537 | -33.9946 | -2.0686
1.1598 | -33.9948 | -2.0847
1.1659 | -33.9947 | -2.0997
1.172 | -33.9946 | -2.1152
1.1781 | -33.9945 | -2.1302
1.1842 | -33.9945 | -2.1455
1.1903 | -33.9948 | -2.158
1.1964 | -33.9948 | -2.1722
1.2025 | -33.995 | -2.1874
1.2086 | -33.9954 | -2.2012
1.2147 | -33.9954 | -2.2151
1.2209 | -33.9958 | -2.2274
1.227 | -33.9965 | -2.2395
1.2331 | -33.9969 | -2.2527
1.2392 | -33.9973 | -2.2654
1.2453 | -33.9978 | -2.2779
1.2514 | -33.9982 | -2.2889
1.2575 | -33.9988 | -2.3011
1.2636 | -33.9995 | -2.3111
1.2697 | -34.0001 | -2.3213
1.2758 | -34.0007 | -2.3337
1.2819 | -34.0014 | -2.3444
1.288 | -34.0021 | -2.356
1.2941 | -34.0028 | -2.3658
1.3002 | -34.0034 | -2.3746
1.3063 | -34.0039 | -2.3847
1.3124 | -34.0045 | -2.3952
1.3185 | -34.0051 | -2.4046
1.3246 | -34.0056 | -2.4143
1.3307 | -34.0064 | -2.4238
1.3368 | -34.0071 | -2.4323
1.3429 | -34.0076 | -2.4421
1.349 | -34.0081 | -2.4518
1.3551 | -34.0087 | -2.4607
1.3613 | -34.0092 | -2.4691
1.3674 | -34.01 | -2.4791
1.3735 | -34.0105 | -2.489
1.3796 | -34.011 | -2.4983
1.3857 | -34.0114 | -2.5078
1.3918 | -34.012 | -2.518
1.3979 | -34.0126 | -2.5272
1.404 | -34.0132 | -2.5386
1.4101 | -34.0136 | -2.5487
1.4162 | -34.014 | -2.5593
1.4223 | -34.0146 | -2.5699
1.4284 | -34.0151 | -2.5798
1.4345 | -34.0155 | -2.5911
1.4406 | -34.0159 | -2.604
1.4467 | -34.0164 | -2.6155
1.4528 | -34.017 | -2.6267
1.4589 | -34.0176 | -2.6387
1.465 | -34.0179 | -2.6492
1.4711 | -34.0184 | -2.6602
1.4772 | -34.019 | -2.6738
1.4833 | -34.0195 | -2.6865
1.4894 | -34.0201 | -2.6988
1.4955 | -34.0207 | -2.7112
1.5016 | -34.0213 | -2.7238
1.5078 | -34.0217 | -2.7381
1.5139 | -34.0221 | -2.751
1.52 | -34.0225 | -2.7636
1.5261 | -34.0228 | -2.7764
1.5322 | -34.0231 | -2.791
1.5383 | -34.0235 | -2.8041
1.5444 | -34.0239 | -2.8179
1.5505 | -34.0243 | -2.832
1.5566 | -34.0248 | -2.8457
1.5627 | -34.0252 | -2.8599
1.5688 | -34.0258 | -2.8742
1.5749 | -34.026 | -2.8874
1.581 | -34.0262 | -2.9024
1.5871 | -34.0265 | -2.9165
1.5932 | -34.0268 | -2.93
1.5993 | -34.0271 | -2.9449
1.6054 | -34.0274 | -2.9589
1.6115 | -34.0277 | -2.9738
1.6176 | -34.0278 | -2.9876
1.6237 | -34.0279 | -3.002
1.6298 | -34.0281 | -3.0158
1.6359 | -34.0283 | -3.0293
1.642 | -34.0285 | -3.0422
1.6482 | -34.0286 | -3.0561
1.6543 | -34.0289 | -3.0704
1.6604 | -34.0291 | -3.0835
1.6665 | -34.0291 | -3.0969
1.6726 | -34.0292 | -3.1106
1.6787 | -34.0295 | -3.124
1.6848 | -34.0293 | -3.1373
1.6909 | -34.0295 | -3.1509
1.697 | -34.0297 | -3.1641
1.7031 | -34.0297 | -3.1775
1.7092 | -34.0297 | -3.1903
1.7153 | -34.0298 | -3.2034
1.7214 | -34.0296 | -3.2162
1.7275 | -34.0295 | -3.2292
1.7336 | -34.0294 | -3.2429
1.7397 | -34.0296 | -3.2554
1.7458 | -34.0296 | -3.2686
1.7519 | -34.0295 | -3.2815
1.758 | -34.0297 | -3.2936
1.7641 | -34.0298 | -3.3056
1.7702 | -34.0298 | -3.319
1.7763 | -34.0298 | -3.3313
1.7824 | -34.0297 | -3.3441
1.7885 | -34.0299 | -3.3564
1.7947 | -34.0301 | -3.3691
1.8008 | -34.0301 | -3.3805
1.8069 | -34.0301 | -3.3923
1.813 | -34.0303 | -3.4044
1.8191 | -34.0303 | -3.4172
1.8252 | -34.0306 | -3.4293
1.8313 | -34.0308 | -3.4413
1.8374 | -34.0309 | -3.4533
1.8435 | -34.031 | -3.465
1.8496 | -34.0312 | -3.4779
1.8557 | -34.0315 | -3.4902
1.8618 | -34.0319 | -3.5016
1.8679 | -34.0321 | -3.5132
1.874 | -34.0323 | -3.5246
1.8801 | -34.0326 | -3.5364
1.8862 | -34.0329 | -3.5483
1.8923 | -34.0332 | -3.56
1.8984 | -34.0334 | -3.5714
1.9045 | -34.0336 | -3.5828
1.9106 | -34.0338 | -3.5944
1.9167 | -34.0339 | -3.6061
1.9228 | -34.0342 | -3.6172
1.9289 | -34.0344 | -3.6281
1.9351 | -34.0347 | -3.6384
1.9412 | -34.035 | -3.6492
1.9473 | -34.0353 | -3.6607
1.9534 | -34.0356 | -3.6712
1.9595 | -34.0359 | -3.6823
1.9656 | -34.0362 | -3.6938
1.9717 | -34.0365 | -3.7045
1.9778 | -34.0369 | -3.7151
1.9839 | -34.0373 | -3.7257
1.99 | -34.0375 | -3.736
1.9961 | -34.0378 | -3.7472
2.0022 | -34.038 | -3.7578
2.0083 | -34.0382 | -3.7684
2.0144 | -34.0385 | -3.7785
2.0205 | -34.0389 | -3.7893
2.0266 | -34.0393 | -3.7997
2.0327 | -34.0396 | -3.8111
2.0388 | -34.04 | -3.8222
2.0449 | -34.0402 | -3.8328
2.051 | -34.0405 | -3.8435
2.0571 | -34.0409 | -3.8542
2.0632 | -34.0412 | -3.8652
2.0693 | -34.0414 | -3.8764
2.0754 | -34.0417 | -3.8873
2.0816 | -34.042 | -3.8988
2.0877 | -34.0421 | -3.9096
2.0938 | -34.0426 | -3.9212
2.0999 | -34.0428 | -3.9324
2.106 | -34.0429 | -3.9438
2.1121 | -34.0431 | -3.9551
2.1182 | -34.0434 | -3.9654
2.1243 | -34.0437 | -3.9756
2.1304 | -34.0439 | -3.9866
2.1365 | -34.0442 | -3.9978
2.1426 | -34.0443 | -4.0093
2.1487 | -34.0446 | -4.0199
2.1548 | -34.045 | -4.0313
2.1609 | -34.0454 | -4.0423
2.167 | -34.0458 | -4.0532
2.1731 | -34.0461 | -4.0651
2.1792 | -34.0463 | -4.076
2.1853 | -34.0467 | -4.0878
2.1914 | -34.047 | -4.0983
2.1975 | -34.0474 | -4.1092
2.2036 | -34.0476 | -4.1209
2.2097 | -34.0478 | -4.132
2.2158 | -34.048 | -4.1423
2.222 | -34.0483 | -4.1529
2.2281 | -34.0485 | -4.1639
2.2342 | -34.0488 | -4.1752
2.2403 | -34.0489 | -4.1853
2.2464 | -34.0492 | -4.1957
2.2525 | -34.0495 | -4.207
2.2586 | -34.0498 | -4.217
2.2647 | -34.05 | -4.2273
2.2708 | -34.0502 | -4.2382
2.2769 | -34.0505 | -4.249
2.283 | -34.0506 | -4.2599
2.2891 | -34.0509 | -4.2707
2.2952 | -34.0511 | -4.2811
2.3013 | -34.0513 | -4.2917
2.3074 | -34.0515 | -4.3018
2.3135 | -34.0517 | -4.313
2.3196 | -34.0519 | -4.3232
2.3257 | -34.0521 | -4.3335
2.3318 | -34.0522 | -4.3442
2.3379 | -34.0524 | -4.3549
2.344 | -34.0525 | -4.366
2.3501 | -34.0527 | -4.3767
2.3562 | -34.0528 | -4.387
2.3623 | -34.0529 | -4.3974
2.3685 | -34.053 | -4.4079
2.3746 | -34.0531 | -4.4189
2.3807 | -34.0534 | -4.4293
2.3868 | -34.0535 | -4.4401
2.3929 | -34.0537 | -4.4504
2.399 | -34.0537 | -4.4609
2.4051 | -34.054 | -4.4712
2.4112 | -34.0542 | -4.4817
2.4173 | -34.0543 | -4.4922
2.4234 | -34.0544 | -4.5027
2.4295 | -34.0546 | -4.5134
2.4356 | -34.0547 | -4.5241
2.4417 | -34.0548 | -4.5353
2.4478 | -34.0549 | -4.5461
2.4539 | -34.0551 | -4.5563
2.46 | -34.0552 | -4.5676
2.4661 | -34.0553 | -4.5785
2.4722 | -34.0555 | -4.5886
2.4783 | -34.0555 | -4.5988
2.4844 | -34.0556 | -4.6101
2.4905 | -34.0557 | -4.6207
2.4966 | -34.0559 | -4.6313
2.5027 | -34.056 | -4.6417
2.5089 | -34.0561 | -4.6526
2.515 | -34.0561 | -4.6634
2.5211 | -34.0562 | -4.6742
2.5272 | -34.0563 | -4.6851
2.5333 | -34.0564 | -4.6959
2.5394 | -34.0566 | -4.7066
2.5455 | -34.0568 | -4.7178
2.5516 | -34.0569 | -4.7283
2.5577 | -34.0571 | -4.7396
2.5638 | -34.0573 | -4.7513
2.5699 | -34.0573 | -4.7625
2.576 | -34.0574 | -4.7741
2.5821 | -34.0575 | -4.7852
2.5882 | -34.0577 | -4.7964
2.5943 | -34.0578 | -4.8071
2.6004 | -34.0581 | -4.8179
2.6065 | -34.0582 | -4.8284
2.6126 | -34.0583 | -4.8391
2.6187 | -34.0585 | -4.8497
2.6248 | -34.0587 | -4.8605
2.6309 | -34.0589 | -4.8721
2.637 | -34.0592 | -4.8837
2.6431 | -34.0593 | -4.8941
2.6492 | -34.0595 | -4.9055
2.6554 | -34.0597 | -4.9164
2.6615 | -34.0598 | -4.9274
2.6676 | -34.0601 | -4.9389
2.6737 | -34.0604 | -4.9502
2.6798 | -34.0607 | -4.9613
2.6859 | -34.0609 | -4.9725
2.692 | -34.0611 | -4.9833
2.6981 | -34.0614 | -4.9945
2.7042 | -34.0616 | -5.0057
2.7103 | -34.0619 | -5.0162
2.7164 | -34.0622 | -5.0272
2.7225 | -34.0625 | -5.0385
2.7286 | -34.0627 | -5.0497
2.7347 | -34.063 | -5.0603
2.7408 | -34.0633 | -5.0706
2.7469 | -34.0636 | -5.0819
2.753 | -34.0639 | -5.0926
2.7591 | -34.0642 | -5.1039
2.7652 | -34.0645 | -5.1155
2.7713 | -34.0646 | -5.1263
2.7774 | -34.065 | -5.1374
2.7835 | -34.0652 | -5.1485
2.7896 | -34.0654 | -5.159
2.7958 | -34.0657 | -5.1697
2.8019 | -34.066 | -5.1802
2.808 | -34.0663 | -5.191
2.8141 | -34.0666 | -5.2022
2.8202 | -34.067 | -5.2133
2.8263 | -34.0673 | -5.2243
2.8324 | -34.0677 | -5.2358
2.8385 | -34.0681 | -5.2468
2.8446 | -34.0685 | -5.2578
2.8507 | -34.0689 | -5.2687
2.8568 | -34.0693 | -5.2794
2.8629 | -34.0696 | -5.2902
2.869 | -34.0699 | -5.3013
2.8751 | -34.0703 | -5.3123
2.8812 | -34.0707 | -5.3238
2.8873 | -34.0711 | -5.335
2.8934 | -34.0714 | -5.346
2.8995 | -34.0718 | -5.357
2.9056 | -34.0721 | -5.3682
2.9117 | -34.0724 | -5.3793
2.9178 | -34.0728 | -5.3902
2.9239 | -34.0732 | -5.4019
2.93 | -34.0737 | -5.4134
2.9361 | -34.074 | -5.4245
2.9423 | -34.0743 | -5.4352
2.9484 | -34.0747 | -5.4465
2.9545 | -34.0751 | -5.4576
2.9606 | -34.0755 | -5.469
2.9667 | -34.0759 | -5.4801
2.9728 | -34.0763 | -5.4915
2.9789 | -34.0767 | -5.5027
2.985 | -34.0771 | -5.514
2.9911 | -34.0775 | -5.5253
2.9972 | -34.078 | -5.5364
3.0033 | -34.0784 | -5.5478
3.0094 | -34.0788 | -5.559
3.0155 | -34.0793 | -5.5702
3.0216 | -34.0797 | -5.5813
3.0277 | -34.08 | -5.5931
3.0338 | -34.0804 | -5.6045
3.0399 | -34.0809 | -5.6157
3.046 | -34.0812 | -5.6273
3.0521 | -34.0816 | -5.6386
3.0582 | -34.0819 | -5.6503
3.0643 | -34.0823 | -5.6617
3.0704 | -34.0826 | -5.6736
3.0765 | -34.0829 | -5.6847
3.0827 | -34.0832 | -5.6959
3.0888 | -34.0836 | -5.7071
3.0949 | -34.084 | -5.7183
3.101 | -34.0843 | -5.7296
3.1071 | -34.0847 | -5.741
3.1132 | -34.0851 | -5.7528
3.1193 | -34.0854 | -5.7641
3.1254 | -34.0858 | -5.7753
3.1315 | -34.0862 | -5.7873
3.1376 | -34.0865 | -5.7987
3.1437 | -34.0869 | -5.81
3.1498 | -34.0873 | -5.8218
3.1559 | -34.0876 | -5.8333
3.162 | -34.0879 | -5.8451
3.1681 | -34.0882 | -5.8568
3.1742 | -34.0885 | -5.8681
3.1803 | -34.0888 | -5.8797
3.1864 | -34.0892 | -5.8907
3.1925 | -34.0895 | -5.9024
3.1986 | -34.0898 | -5.914
3.2047 | -34.0901 | -5.9259
3.2108 | -34.0904 | -5.9369
3.2169 | -34.0907 | -5.9483
3.223 | -34.0909 | -5.9597
3.2292 | -34.0912 | -5.9713
3.2353 | -34.0914 | -5.9821
3.2414 | -34.0917 | -5.9934
3.2475 | -34.0919 | -6.0046
3.2536 | -34.0922 | -6.0158
3.2597 | -34.0925 | -6.0266
3.2658 | -34.0927 | -6.0374
3.2719 | -34.093 | -6.0482
3.278 | -34.0932 | -6.0596
3.2841 | -34.0935 | -6.0706
3.2902 | -34.0938 | -6.0819
3.2963 | -34.094 | -6.0931
3.3024 | -34.0941 | -6.104
3.3085 | -34.0944 | -6.1154
3.3146 | -34.0946 | -6.1265
3.3207 | -34.0949 | -6.1378
3.3268 | -34.0951 | -6.1492
3.3329 | -34.0953 | -6.1598
3.339 | -34.0955 | -6.1705
3.3451 | -34.0957 | -6.1816
3.3512 | -34.0959 | -6.1922
3.3573 | -34.0962 | -6.2034
3.3634 | -34.0964 | -6.2144
3.3696 | -34.0966 | -6.2256
3.3757 | -34.0967 | -6.2364
3.3818 | -34.097 | -6.2471
3.3879 | -34.0972 | -6.2582
3.394 | -34.0973 | -6.2694
3.4001 | -34.0976 | -6.2803
3.4062 | -34.0977 | -6.291
3.4123 | -34.0979 | -6.3016
3.4184 | -34.0981 | -6.3129
3.4245 | -34.0983 | -6.3236
3.4306 | -34.0984 | -6.3344
3.4367 | -34.0986 | -6.3449
3.4428 | -34.0988 | -6.3553
3.4489 | -34.099 | -6.3663
3.455 | -34.0991 | -6.3772
3.4611 | -34.0993 | -6.3885
3.4672 | -34.0995 | -6.3995
3.4733 | -34.0997 | -6.4106
3.4794 | -34.0999 | -6.4214
3.4855 | -34.1001 | -6.4318
3.4916 | -34.1003 | -6.4423
3.4977 | -34.1004 | -6.4535
3.5038 | -34.1006 | -6.4644
3.5099 | -34.1008 | -6.4748
3.5161 | -34.1009 | -6.4853
3.5222 | -34.1011 | -6.4959
3.5283 | -34.1012 | -6.5068
3.5344 | -34.1014 | -6.5176
3.5405 | -34.1016 | -6.5283
3.5466 | -34.1018 | -6.5393
3.5527 | -34.1019 | -6.5503
3.5588 | -34.102 | -6.5611
3.5649 | -34.1021 | -6.5719
3.571 | -34.1023 | -6.5826
3.5771 | -34.1024 | -6.5937
3.5832 | -34.1026 | -6.6039
3.5893 | -34.1028 | -6.6146
3.5954 | -34.103 | -6.6255
3.6015 | -34.1032 | -6.6365
3.6076 | -34.1034 | -6.6472
3.6137 | -34.1036 | -6.6575
3.6198 | -34.1039 | -6.668
3.6259 | -34.104 | -6.6791
3.632 | -34.1042 | -6.6898
3.6381 | -34.1044 | -6.7007
3.6442 | -34.1046 | -6.7114
3.6503 | -34.1048 | -6.7218
3.6565 | -34.105 | -6.7328
3.6626 | -34.1052 | -6.7436
3.6687 | -34.1054 | -6.7545
3.6748 | -34.1057 | -6.7651
3.6809 | -34.106 | -6.776
3.687 | -34.1061 | -6.7871
3.6931 | -34.1064 | -6.7976
3.6992 | -34.1066 | -6.8085
3.7053 | -34.1069 | -6.8193
3.7114 | -34.107 | -6.8302
3.7175 | -34.1073 | -6.8411
3.7236 | -34.1075 | -6.8515
3.7297 | -34.1077 | -6.8623
3.7358 | -34.108 | -6.8733
3.7419 | -34.1082 | -6.8843
3.748 | -34.1084 | -6.895
3.7541 | -34.1086 | -6.9058
3.7602 | -34.1089 | -6.9166
3.7663 | -34.1091 | -6.9273
3.7724 | -34.1094 | -6.9382
3.7785 | -34.1097 | -6.9488
3.7846 | -34.1099 | -6.9596
3.7907 | -34.1102 | -6.9706
3.7969 | -34.1104 | -6.9815
3.803 | -34.1107 | -6.992
3.8091 | -34.1109 | -7.0027
3.8152 | -34.1112 | -7.0134
3.8213 | -34.1114 | -7.0245
3.8274 | -34.1117 | -7.0351
3.8335 | -34.112 | -7.0461
3.8396 | -34.1123 | -7.0569
3.8457 | -34.1126 | -7.0678
3.8518 | -34.113 | -7.0786
3.8579 | -34.1133 | -7.0894
3.864 | -34.1135 | -7.1003
3.8701 | -34.1139 | -7.1116
3.8762 | -34.1141 | -7.1222
3.8823 | -34.1143 | -7.1332
3.8884 | -34.1147 | -7.144
3.8945 | -34.115 | -7.1551
3.9006 | -34.1153 | -7.1662
3.9067 | -34.1155 | -7.1775
3.9128 | -34.1158 | -7.1884
3.9189 | -34.1161 | -7.1991
3.925 | -34.1163 | -7.2101
3.9311 | -34.1166 | -7.2205
3.9372 | -34.1169 | -7.2315
3.9434 | -34.1172 | -7.2424
3.9495 | -34.1176 | -7.2533
3.9556 | -34.1179 | -7.2639
3.9617 | -34.1182 | -7.2744
3.9678 | -34.1185 | -7.285
3.9739 | -34.1188 | -7.2957
3.98 | -34.119 | -7.3067
3.9861 | -34.1193 | -7.3175
3.9922 | -34.1195 | -7.3284
3.9983 | -34.1198 | -7.3389
4.0044 | -34.1201 | -7.3493
4.0105 | -34.1204 | -7.3601
4.0166 | -34.1207 | -7.371
4.0227 | -34.121 | -7.3817
4.0288 | -34.1212 | -7.3928
4.0349 | -34.1214 | -7.4033
4.041 | -34.1217 | -7.414
4.0471 | -34.1219 | -7.4249
4.0532 | -34.1222 | -7.4353
4.0593 | -34.1225 | -7.4461
4.0654 | -34.1228 | -7.4567
4.0715 | -34.123 | -7.4678
4.0776 | -34.1233 | -7.4784
4.0838 | -34.1236 | -7.4895
4.0899 | -34.1239 | -7.5
4.096 | -34.1242 | -7.5107
4.1021 | -34.1244 | -7.5212
4.1082 | -34.1246 | -7.5316
4.1143 | -34.1249 | -7.5419
4.1204 | -34.1252 | -7.5524
4.1265 | -34.1254 | -7.5629
4.1326 | -34.1257 | -7.5733
4.1387 | -34.126 | -7.5839
4.1448 | -34.1262 | -7.5947
4.1509 | -34.1265 | -7.6053
4.157 | -34.1267 | -7.6163
4.1631 | -34.127 | -7.6271
4.1692 | -34.1272 | -7.6376
4.1753 | -34.1275 | -7.6482
4.1814 | -34.1277 | -7.6584
4.1875 | -34.128 | -7.6692
4.1936 | -34.1282 | -7.6799
4.1997 | -34.1285 | -7.6902
4.2058 | -34.1287 | -7.7011
4.2119 | -34.1289 | -7.7115
4.218 | -34.1292 | -7.7218
4.2241 | -34.1294 | -7.7322
4.2303 | -34.1297 | -7.7427
4.2364 | -34.1299 | -7.7532
4.2425 | -34.1301 | -7.7638
4.2486 | -34.1303 | -7.7745
4.2547 | -34.1304 | -7.7848
4.2608 | -34.1306 | -7.7955
4.2669 | -34.1308 | -7.8057
4.273 | -34.1311 | -7.8164
4.2791 | -34.1313 | -7.8269
4.2852 | -34.1315 | -7.8375
4.2913 | -34.1316 | -7.8477
4.2974 | -34.1319 | -7.8584
4.3035 | -34.1321 | -7.8693
4.3096 | -34.1323 | -7.88
4.3157 | -34.1325 | -7.8903
4.3218 | -34.1327 | -7.9007
4.3279 | -34.133 | -7.9112
4.334 | -34.1331 | -7.9216
4.3401 | -34.1333 | -7.9317
4.3462 | -34.1336 | -7.9423
4.3523 | -34.1338 | -7.9524
4.3584 | -34.134 | -7.9626
4.3645 | -34.1342 | -7.973
4.3707 | -34.1344 | -7.9832
4.3768 | -34.1347 | -7.9935
4.3829 | -34.1349 | -8.0042
4.389 | -34.1351 | -8.0146
4.3951 | -34.1353 | -8.0248
4.4012 | -34.1355 | -8.0355
4.4073 | -34.1357 | -8.0458
4.4134 | -34.1359 | -8.0561
4.4195 | -34.1362 | -8.0665
4.4256 | -34.1364 | -8.0768
4.4317 | -34.1367 | -8.0871
4.4378 | -34.1368 | -8.0971
4.4439 | -34.137 | -8.1071
4.45 | -34.1372 | -8.1175
4.4561 | -34.1373 | -8.1278
4.4622 | -34.1376 | -8.1381
4.4683 | -34.1378 | -8.1484
4.4744 | -34.138 | -8.159
4.4805 | -34.1382 | -8.1689
4.4866 | -34.1384 | -8.1793
4.4927 | -34.1387 | -8.1898
4.4988 | -34.1389 | -8.2
4.5049 | -34.1391 | -8.21
4.511 | -34.1394 | -8.22
4.5172 | -34.1396 | -8.2303
4.5233 | -34.1397 | -8.241
4.5294 | -34.14 | -8.2511
4.5355 | -34.1402 | -8.2614
4.5416 | -34.1404 | -8.2717
4.5477 | -34.1407 | -8.2819
4.5538 | -34.1409 | -8.2919
4.5599 | -34.1411 | -8.3021
4.566 | -34.1414 | -8.3124
4.5721 | -34.1416 | -8.3227
4.5782 | -34.1418 | -8.3332
4.5843 | -34.1421 | -8.3436
4.5904 | -34.1423 | -8.3539
4.5965 | -34.1426 | -8.364
4.6026 | -34.1428 | -8.3747
4.6087 | -34.1431 | -8.3848
4.6148 | -34.1433 | -8.3954
4.6209 | -34.1435 | -8.4056
4.627 | -34.1438 | -8.416
4.6331 | -34.144 | -8.4262
4.6392 | -34.1442 | -8.4366
4.6453 | -34.1444 | -8.4468
4.6514 | -34.1447 | -8.4572
4.6576 | -34.1449 | -8.4679
4.6637 | -34.1451 | -8.4778
4.6698 | -34.1453 | -8.4879
4.6759 | -34.1455 | -8.4986
4.682 | -34.1458 | -8.5088
4.6881 | -34.1459 | -8.519
4.6942 | -34.1462 | -8.5292
4.7003 | -34.1464 | -8.5396
4.7064 | -34.1467 | -8.5498
4.7125 | -34.1469 | -8.5604
4.7186 | -34.1471 | -8.5704
4.7247 | -34.1474 | -8.5808
4.7308 | -34.1477 | -8.5911
4.7369 | -34.1479 | -8.6015
4.743 | -34.1482 | -8.6122
4.7491 | -34.1484 | -8.6228
4.7552 | -34.1486 | -8.6332
4.7613 | -34.1489 | -8.6436
4.7674 | -34.1491 | -8.6538
4.7735 | -34.1493 | -8.6645
4.7796 | -34.1496 | -8.6751
4.7857 | -34.1499 | -8.6855
4.7918 | -34.1501 | -8.6956
4.7979 | -34.1504 | -8.7059
4.8041 | -34.1507 | -8.7162
4.8102 | -34.1509 | -8.7263
4.8163 | -34.1512 | -8.7367
4.8224 | -34.1514 | -8.747
4.8285 | -34.1517 | -8.7571
4.8346 | -34.152 | -8.7672
4.8407 | -34.1522 | -8.7777
4.8468 | -34.1525 | -8.7883
4.8529 | -34.1528 | -8.7989
4.859 | -34.153 | -8.8096
4.8651 | -34.1533 | -8.8198
4.8712 | -34.1535 | -8.83
4.8773 | -34.1538 | -8.8405
4.8834 | -34.154 | -8.8508
4.8895 | -34.1543 | -8.8611
4.8956 | -34.1545 | -8.8715
4.9017 | -34.1547 | -8.8818
4.9078 | -34.155 | -8.8924
4.9139 | -34.1553 | -8.9025
4.92 | -34.1556 | -8.9126
4.9261 | -34.1559 | -8.9227
4.9322 | -34.1562 | -8.933
4.9383 | -34.1565 | -8.9433
4.9445 | -34.1567 | -8.9537
4.9506 | -34.157 | -8.9639
4.9567 | -34.1573 | -8.9742
4.9628 | -34.1575 | -8.9845
4.9689 | -34.1578 | -8.9952
4.975 | -34.158 | -9.0059
4.9811 | -34.1583 | -9.0159
4.9872 | -34.1586 | -9.0261
4.9933 | -34.159 | -9.0368
4.9994 | -34.1593 | -9.0475
5.0055 | -34.1595 | -9.0579
5.0116 | -34.1598 | -9.0687
5.0177 | -34.1602 | -9.0793
5.0238 | -34.1605 | -9.09
5.0299 | -34.1608 | -9.1007
5.036 | -34.1611 | -9.1111
5.0421 | -34.1614 | -9.1214
5.0482 | -34.1617 | -9.1318
5.0543 | -34.162 | -9.142
5.0604 | -34.1623 | -9.1524
5.0665 | -34.1626 | -9.1626
5.0726 | -34.1629 | -9.1734
5.0787 | -34.1632 | -9.1839
5.0848 | -34.1635 | -9.1945
5.091 | -34.1638 | -9.2048
5.0971 | -34.1641 | -9.2149
5.1032 | -34.1644 | -9.2257
5.1093 | -34.1647 | -9.2361
5.1154 | -34.165 | -9.2464
5.1215 | -34.1652 | -9.2567
5.1276 | -34.1656 | -9.267
5.1337 | -34.1659 | -9.2772
5.1398 | -34.1662 | -9.2874
5.1459 | -34.1665 | -9.298
5.152 | -34.1668 | -9.3086
5.1581 | -34.1671 | -9.3192
5.1642 | -34.1674 | -9.3297
5.1703 | -34.1677 | -9.3401
5.1764 | -34.1681 | -9.3504
5.1825 | -34.1683 | -9.3606
5.1886 | -34.1687 | -9.3707
5.1947 | -34.1689 | -9.3812
5.2008 | -34.1693 | -9.3917
5.2069 | -34.1696 | -9.4023
5.213 | -34.1699 | -9.4125
5.2191 | -34.1702 | -9.423
5.2252 | -34.1705 | -9.4335
5.2314 | -34.1708 | -9.4441
5.2375 | -34.1711 | -9.4543
5.2436 | -34.1714 | -9.4648
5.2497 | -34.1717 | -9.4749
5.2558 | -34.172 | -9.4857
5.2619 | -34.1723 | -9.4963
5.268 | -34.1727 | -9.5069
5.2741 | -34.173 | -9.5174
5.2802 | -34.1733 | -9.5282
5.2863 | -34.1737 | -9.5386
5.2924 | -34.174 | -9.5492
5.2985 | -34.1743 | -9.5597
5.3046 | -34.1746 | -9.5702
5.3107 | -34.1749 | -9.5808
5.3168 | -34.1752 | -9.5914
5.3229 | -34.1755 | -9.6018
5.329 | -34.1758 | -9.6123
5.3351 | -34.1762 | -9.6225
5.3412 | -34.1765 | -9.6329
5.3473 | -34.1768 | -9.6434
5.3534 | -34.1771 | -9.654
5.3595 | -34.1774 | -9.6645
5.3656 | -34.1778 | -9.6749
5.3717 | -34.1781 | -9.6856
5.3779 | -34.1784 | -9.6964
5.384 | -34.1787 | -9.7068
5.3901 | -34.179 | -9.7176
5.3962 | -34.1793 | -9.7279
5.4023 | -34.1797 | -9.7385
5.4084 | -34.18 | -9.7489
5.4145 | -34.1803 | -9.7592
5.4206 | -34.1806 | -9.7697
5.4267 | -34.181 | -9.7805
5.4328 | -34.1813 | -9.791
5.4389 | -34.1816 | -9.8012
5.445 | -34.1819 | -9.812
5.4511 | -34.1822 | -9.8223
5.4572 | -34.1825 | -9.8329
5.4633 | -34.1829 | -9.8433
5.4694 | -34.1831 | -9.8538
5.4755 | -34.1834 | -9.8641
5.4816 | -34.1838 | -9.8746
5.4877 | -34.1841 | -9.885
5.4938 | -34.1844 | -9.8952
5.4999 | -34.1847 | -9.9056
5.506 | -34.1851 | -9.916
5.5121 | -34.1854 | -9.9264
5.5183 | -34.1857 | -9.9365
5.5244 | -34.186 | -9.9472
5.5305 | -34.1863 | -9.9576
5.5366 | -34.1866 | -9.9678
5.5427 | -34.1869 | -9.9784
5.5488 | -34.1872 | -9.9889
5.5549 | -34.1875 | -9.9993
5.561 | -34.1878 | -10.01
5.5671 | -34.1882 | -10.0202
5.5732 | -34.1885 | -10.0309
5.5793 | -34.1888 | -10.0416
5.5854 | -34.1892 | -10.052
5.5915 | -34.1894 | -10.0627
5.5976 | -34.1897 | -10.0732
5.6037 | -34.1901 | -10.0838
5.6098 | -34.1903 | -10.0942
5.6159 | -34.1907 | -10.1044
5.622 | -34.191 | -10.1149
5.6281 | -34.1913 | -10.1254
5.6342 | -34.1916 | -10.1362
5.6403 | -34.1919 | -10.1467
5.6464 | -34.1922 | -10.1573
5.6525 | -34.1925 | -10.1678
5.6586 | -34.1928 | -10.178
5.6648 | -34.1932 | -10.1885
5.6709 | -34.1935 | -10.1991
5.677 | -34.1938 | -10.2096
5.6831 | -34.1941 | -10.2201
5.6892 | -34.1944 | -10.2308
5.6953 | -34.1947 | -10.2412
5.7014 | -34.1951 | -10.2517
5.7075 | -34.1953 | -10.2622
5.7136 | -34.1956 | -10.2725
5.7197 | -34.1959 | -10.283
5.7258 | -34.1962 | -10.2933
5.7319 | -34.1965 | -10.3037
5.738 | -34.1968 | -10.3142
5.7441 | -34.1972 | -10.3246
5.7502 | -34.1975 | -10.3348
5.7563 | -34.1978 | -10.3452
5.7624 | -34.1982 | -10.3559
5.7685 | -34.1985 | -10.3663
5.7746 | -34.1988 | -10.377
5.7807 | -34.199 | -10.3875
5.7868 | -34.1993 | -10.3982
5.7929 | -34.1996 | -10.4086
5.799 | -34.1999 | -10.4188
5.8052 | -34.2002 | -10.4291
5.8113 | -34.2005 | -10.4394
5.8174 | -34.2008 | -10.4497
5.8235 | -34.2011 | -10.4602
5.8296 | -34.2014 | -10.4707
5.8357 | -34.2017 | -10.481
5.8418 | -34.202 | -10.4916
5.8479 | -34.2023 | -10.5021
5.854 | -34.2026 | -10.5126
5.8601 | -34.2029 | -10.5229
5.8662 | -34.2032 | -10.5335
5.8723 | -34.2035 | -10.5437
5.8784 | -34.2038 | -10.554
5.8845 | -34.2041 | -10.5643
5.8906 | -34.2044 | -10.5748
5.8967 | -34.2047 | -10.585
5.9028 | -34.205 | -10.5953
5.9089 | -34.2054 | -10.6055
5.915 | -34.2057 | -10.6157
5.9211 | -34.206 | -10.6261
5.9272 | -34.2063 | -10.6362
5.9333 | -34.2066 | -10.6465
5.9394 | -34.2069 | -10.6568
5.9455 | -34.2072 | -10.667
5.9517 | -34.2076 | -10.6772
5.9578 | -34.2079 | -10.6877
5.9639 | -34.2082 | -10.6977
5.97 | -34.2086 | -10.7082
5.9761 | -34.2089 | -10.7184
5.9822 | -34.2092 | -10.7289
5.9883 | -34.2095 | -10.7392
5.9944 | -34.2098 | -10.7497
6.0005 | -34.2102 | -10.7602
6.0066 | -34.2105 | -10.7704
6.0127 | -34.2108 | -10.7804
6.0188 | -34.2111 | -10.7909
6.0249 | -34.2114 | -10.8013
6.031 | -34.2117 | -10.8114
6.0371 | -34.2121 | -10.8218
6.0432 | -34.2124 | -10.832
6.0493 | -34.2127 | -10.8422
6.0554 | -34.213 | -10.8524
6.0615 | -34.2134 | -10.8627
6.0676 | -34.2137 | -10.8732
6.0737 | -34.214 | -10.8836
6.0798 | -34.2143 | -10.8939
6.0859 | -34.2147 | -10.9042
6.0921 | -34.215 | -10.9143
6.0982 | -34.2153 | -10.9246
6.1043 | -34.2156 | -10.9348
6.1104 | -34.216 | -10.9453
6.1165 | -34.2163 | -10.9552
6.1226 | -34.2166 | -10.9654
6.1287 | -34.2169 | -10.9756
6.1348 | -34.2173 | -10.9859
6.1409 | -34.2176 | -10.9961
6.147 | -34.2179 | -11.0063
6.1531 | -34.2182 | -11.0166
6.1592 | -34.2185 | -11.0267
6.1653 | -34.2188 | -11.0369
6.1714 | -34.2191 | -11.0472
6.1775 | -34.2195 | -11.0576
6.1836 | -34.2198 | -11.0679
6.1897 | -34.2201 | -11.0781
6.1958 | -34.2204 | -11.0886
6.2019 | -34.2208 | -11.099
6.208 | -34.2211 | -11.1092
6.2141 | -34.2214 | -11.1195
6.2202 | -34.2218 | -11.1298
6.2263 | -34.2221 | -11.14
6.2325 | -34.2224 | -11.15
6.2386 | -34.2228 | -11.16
6.2447 | -34.2231 | -11.1704
6.2508 | -34.2234 | -11.1807
6.2569 | -34.2238 | -11.1909
6.263 | -34.2241 | -11.2011
6.2691 | -34.2244 | -11.2113
6.2752 | -34.2247 | -11.2215
6.2813 | -34.225 | -11.2317
6.2874 | -34.2253 | -11.2419
6.2935 | -34.2256 | -11.2521
6.2996 | -34.2259 | -11.2622
6.3057 | -34.2263 | -11.2726
6.3118 | -34.2266 | -11.2829
6.3179 | -34.2269 | -11.2931
6.324 | -34.2272 | -11.3034
6.3301 | -34.2275 | -11.3136
6.3362 | -34.2279 | -11.3238
6.3423 | -34.2282 | -11.3342
6.3484 | -34.2286 | -11.3445
6.3545 | -34.2289 | -11.3545
6.3606 | -34.2292 | -11.3645
6.3667 | -34.2295 | -11.3746
6.3728 | -34.2298 | -11.3848
6.379 | -34.2301 | -11.395
6.3851 | -34.2305 | -11.405
6.3912 | -34.2307 | -11.4152
6.3973 | -34.2311 | -11.4254
6.4034 | -34.2314 | -11.4355
6.4095 | -34.2317 | -11.4457
6.4156 | -34.232 | -11.4559
6.4217 | -34.2323 | -11.4661
6.4278 | -34.2327 | -11.4763
6.4339 | -34.233 | -11.4865
6.44 | -34.2334 | -11.4966
6.4461 | -34.2337 | -11.5069
6.4522 | -34.234 | -11.5169
6.4583 | -34.2343 | -11.5272
6.4644 | -34.2346 | -11.5375
6.4705 | -34.2349 | -11.5474
6.4766 | -34.2352 | -11.5576
6.4827 | -34.2356 | -11.5678
6.4888 | -34.2358 | -11.5779
6.4949 | -34.2362 | -11.5878
6.501 | -34.2365 | -11.5981
6.5071 | -34.2368 | -11.6081
6.5132 | -34.2372 | -11.6183
6.5194 | -34.2375 | -11.6285
6.5255 | -34.2378 | -11.6385
6.5316 | -34.2381 | -11.6488
6.5377 | -34.2384 | -11.659
6.5438 | -34.2387 | -11.669
6.5499 | -34.2391 | -11.6791
6.556 | -34.2394 | -11.6892
6.5621 | -34.2397 | -11.6996
6.5682 | -34.24 | -11.7098
6.5743 | -34.2404 | -11.7199
6.5804 | -34.2407 | -11.7302
6.5865 | -34.241 | -11.7404
6.5926 | -34.2414 | -11.7504
6.5987 | -34.2417 | -11.7603
6.6048 | -34.242 | -11.7706
6.6109 | -34.2424 | -11.7808
6.617 | -34.2427 | -11.7911
6.6231 | -34.243 | -11.8012
6.6292 | -34.2433 | -11.8114
6.6353 | -34.2436 | -11.8218
6.6414 | -34.244 | -11.8319
6.6475 | -34.2443 | -11.8421
6.6536 | -34.2446 | -11.8522
6.6597 | -34.2449 | -11.8621
6.6659 | -34.2453 | -11.8723
6.672 | -34.2456 | -11.8823
6.6781 | -34.2459 | -11.8925
6.6842 | -34.2463 | -11.9025
6.6903 | -34.2466 | -11.9126
6.6964 | -34.247 | -11.9227
6.7025 | -34.2473 | -11.9327
6.7086 | -34.2476 | -11.9428
6.7147 | -34.2479 | -11.953
6.7208 | -34.2482 | -11.9632
6.7269 | -34.2485 | -11.9733
6.733 | -34.2488 | -11.9832
6.7391 | -34.2492 | -11.9934
6.7452 | -34.2495 | -12.0033
6.7513 | -34.2499 | -12.0132
6.7574 | -34.2502 | -12.0234
6.7635 | -34.2505 | -12.0334
6.7696 | -34.2509 | -12.0434
6.7757 | -34.2512 | -12.0535
6.7818 | -34.2515 | -12.0637
6.7879 | -34.2519 | -12.0737
6.794 | -34.2522 | -12.0836
6.8001 | -34.2525 | -12.0937
6.8063 | -34.2528 | -12.1038
6.8124 | -34.2532 | -12.1141
6.8185 | -34.2535 | -12.1244
6.8246 | -34.2538 | -12.1344
6.8307 | -34.2541 | -12.1444
6.8368 | -34.2544 | -12.1547
6.8429 | -34.2548 | -12.1649
6.849 | -34.2551 | -12.175
6.8551 | -34.2555 | -12.185
6.8612 | -34.2558 | -12.1951
6.8673 | -34.2561 | -12.2052
6.8734 | -34.2564 | -12.2154
6.8795 | -34.2568 | -12.2255
6.8856 | -34.2571 | -12.2356
6.8917 | -34.2574 | -12.2457
6.8978 | -34.2578 | -12.2556
6.9039 | -34.2581 | -12.2654
6.91 | -34.2585 | -12.2756
6.9161 | -34.2588 | -12.2855
6.9222 | -34.2591 | -12.2956
6.9283 | -34.2595 | -12.3054
6.9344 | -34.2598 | -12.3155
6.9405 | -34.2602 | -12.3255
6.9466 | -34.2605 | -12.3358
6.9528 | -34.2609 | -12.346
6.9589 | -34.2612 | -12.3562
6.965 | -34.2616 | -12.3663
6.9711 | -34.2619 | -12.3765
6.9772 | -34.2623 | -12.3865
6.9833 | -34.2626 | -12.3966
6.9894 | -34.2629 | -12.4068
6.9955 | -34.2633 | -12.4167
7.0016 | -34.2636 | -12.4267
7.0077 | -34.264 | -12.4368
7.0138 | -34.2643 | -12.4468
7.0199 | -34.2646 | -12.4568
7.026 | -34.265 | -12.4668
7.0321 | -34.2653 | -12.4767
7.0382 | -34.2656 | -12.4865
7.0443 | -34.266 | -12.4967
7.0504 | -34.2663 | -12.5068
7.0565 | -34.2667 | -12.5168
7.0626 | -34.267 | -12.5267
7.0687 | -34.2674 | -12.5364
7.0748 | -34.2677 | -12.5462
7.0809 | -34.2681 | -12.5562
7.087 | -34.2684 | -12.5663
7.0932 | -34.2687 | -12.5764
7.0993 | -34.2691 | -12.5865
7.1054 | -34.2694 | -12.5964
7.1115 | -34.2698 | -12.6062
7.1176 | -34.2701 | -12.6161
7.1237 | -34.2704 | -12.6261
7.1298 | -34.2708 | -12.636
7.1359 | -34.2711 | -12.6461
7.142 | -34.2715 | -12.6558
7.1481 | -34.2718 | -12.6657
7.1542 | -34.2721 | -12.6757
7.1603 | -34.2725 | -12.6857
7.1664 | -34.2728 | -12.6957
7.1725 | -34.2732 | -12.7057
7.1786 | -34.2736 | -12.7156
7.1847 | -34.2739 | -12.7257
7.1908 | -34.2742 | -12.7359
7.1969 | -34.2745 | -12.746
7.203 | -34.2749 | -12.7558
7.2091 | -34.2753 | -12.7656
7.2152 | -34.2756 | -12.7756
7.2213 | -34.2759 | -12.7854
7.2274 | -34.2763 | -12.7954
7.2335 | -34.2766 | -12.8053
7.2397 | -34.277 | -12.8152
7.2458 | -34.2774 | -12.8249
7.2519 | -34.2777 | -12.8349
7.258 | -34.2781 | -12.845
7.2641 | -34.2784 | -12.855
7.2702 | -34.2787 | -12.8651
7.2763 | -34.2791 | -12.8751
7.2824 | -34.2795 | -12.8851
7.2885 | -34.2798 | -12.8951
7.2946 | -34.2802 | -12.905
7.3007 | -34.2805 | -12.915
7.3068 | -34.2808 | -12.9252
7.3129 | -34.2812 | -12.9352
7.319 | -34.2815 | -12.9451
7.3251 | -34.2819 | -12.9549
7.3312 | -34.2822 | -12.9649
7.3373 | -34.2826 | -12.9746
7.3434 | -34.2829 | -12.9845
7.3495 | -34.2833 | -12.9942
7.3556 | -34.2837 | -13.0043
7.3617 | -34.284 | -13.0144
7.3678 | -34.2844 | -13.0242
7.3739 | -34.2847 | -13.0343
7.3801 | -34.2851 | -13.0445
7.3862 | -34.2854 | -13.0544
7.3923 | -34.2858 | -13.0645
7.3984 | -34.2861 | -13.0742
7.4045 | -34.2865 | -13.0841
7.4106 | -34.2869 | -13.094
7.4167 | -34.2872 | -13.104
7.4228 | -34.2876 | -13.114
7.4289 | -34.2879 | -13.1239
7.435 | -34.2883 | -13.1337
7.4411 | -34.2886 | -13.1436
7.4472 | -34.289 | -13.1536
7.4533 | -34.2893 | -13.1633
7.4594 | -34.2897 | -13.1735
7.4655 | -34.2901 | -13.1833
7.4716 | -34.2904 | -13.1934
7.4777 | -34.2908 | -13.2033
7.4838 | -34.2911 | -13.213
7.4899 | -34.2915 | -13.2228
7.496 | -34.2918 | -13.2325
7.5021 | -34.2922 | -13.2424
7.5082 | -34.2925 | -13.2524
7.5143 | -34.2929 | -13.2623
7.5204 | -34.2932 | -13.2723
7.5266 | -34.2936 | -13.2821
7.5327 | -34.2939 | -13.2919
7.5388 | -34.2943 | -13.3021
7.5449 | -34.2946 | -13.3121
7.551 | -34.295 | -13.3219
7.5571 | -34.2954 | -13.3319
7.5632 | -34.2957 | -13.3419
7.5693 | -34.2961 | -13.3517
7.5754 | -34.2965 | -13.3614
7.5815 | -34.2968 | -13.3714
7.5876 | -34.2972 | -13.3815
7.5937 | -34.2975 | -13.3915
7.5998 | -34.2979 | -13.4014
7.6059 | -34.2983 | -13.4115
7.612 | -34.2986 | -13.4214
7.6181 | -34.299 | -13.4312
7.6242 | -34.2994 | -13.4414
7.6303 | -34.2997 | -13.4514
7.6364 | -34.3001 | -13.4614
7.6425 | -34.3004 | -13.4714
7.6486 | -34.3008 | -13.4814
7.6547 | -34.3011 | -13.4912
7.6608 | -34.3015 | -13.5011
7.667 | -34.3018 | -13.511
7.6731 | -34.3022 | -13.521
7.6792 | -34.3026 | -13.5312
7.6853 | -34.3029 | -13.5412
7.6914 | -34.3033 | -13.5512
7.6975 | -34.3037 | -13.5612
7.7036 | -34.304 | -13.5712
7.7097 | -34.3044 | -13.5812
7.7158 | -34.3048 | -13.5912
7.7219 | -34.3051 | -13.601
7.728 | -34.3055 | -13.611
7.7341 | -34.3059 | -13.621
7.7402 | -34.3062 | -13.631
7.7463 | -34.3066 | -13.641
7.7524 | -34.3069 | -13.6511
7.7585 | -34.3073 | -13.6611
7.7646 | -34.3077 | -13.671
7.7707 | -34.308 | -13.6811
7.7768 | -34.3083 | -13.6911
7.7829 | -34.3087 | -13.7013
7.789 | -34.3091 | -13.7115
7.7951 | -34.3094 | -13.7215
7.8012 | -34.3098 | -13.7314
7.8073 | -34.3102 | -13.7416
7.8135 | -34.3106 | -13.7515
7.8196 | -34.3109 | -13.7616
7.8257 | -34.3113 | -13.7718
7.8318 | -34.3116 | -13.7818
7.8379 | -34.312 | -13.7917
7.844 | -34.3124 | -13.8016
7.8501 | -34.3127 | -13.8118
7.8562 | -34.3131 | -13.822
7.8623 | -34.3135 | -13.8318
7.8684 | -34.3138 | -13.8419
7.8745 | -34.3142 | -13.8519
7.8806 | -34.3145 | -13.862
7.8867 | -34.3149 | -13.8718
7.8928 | -34.3153 | -13.8816
7.8989 | -34.3156 | -13.8916
7.905 | -34.316 | -13.9017
7.9111 | -34.3163 | -13.9117
7.9172 | -34.3167 | -13.9217
7.9233 | -34.317 | -13.9318
7.9294 | -34.3174 | -13.9418
7.9355 | -34.3178 | -13.9516
7.9416 | -34.3181 | -13.9617
7.9477 | -34.3185 | -13.9715
7.9539 | -34.3188 | -13.9814
7.96 | -34.3192 | -13.9915
7.9661 | -34.3196 | -14.0015
7.9722 | -34.3199 | -14.0113
7.9783 | -34.3203 | -14.0215
7.9844 | -34.3207 | -14.0314
7.9905 | -34.321 | -14.0415
7.9966 | -34.3214 | -14.0516
8.0027 | -34.3217 | -14.0616
8.0088 | -34.3221 | -14.0716
8.0149 | -34.3224 | -14.0816
8.021 | -34.3228 | -14.0916
8.0271 | -34.3232 | -14.1018
8.0332 | -34.3235 | -14.1117
8.0393 | -34.3239 | -14.1218
8.0454 | -34.3243 | -14.132
8.0515 | -34.3247 | -14.142
8.0576 | -34.325 | -14.1522
8.0637 | -34.3254 | -14.1621
8.0698 | -34.3257 | -14.172
8.0759 | -34.3261 | -14.1819
8.082 | -34.3264 | -14.1918
8.0881 | -34.3268 | -14.2018
8.0942 | -34.3272 | -14.2119
8.1004 | -34.3276 | -14.2219
8.1065 | -34.3279 | -14.2318
8.1126 | -34.3283 | -14.2419
8.1187 | -34.3286 | -14.252
8.1248 | -34.329 | -14.2617
8.1309 | -34.3294 | -14.2716
8.137 | -34.3297 | -14.2814
8.1431 | -34.3301 | -14.2916
8.1492 | -34.3304 | -14.3016
8.1553 | -34.3308 | -14.3117
8.1614 | -34.3312 | -14.3218
8.1675 | -34.3315 | -14.3318
8.1736 | -34.3319 | -14.3419
8.1797 | -34.3322 | -14.3519
8.1858 | -34.3326 | -14.362
8.1919 | -34.333 | -14.3719
8.198 | -34.3333 | -14.3818
8.2041 | -34.3337 | -14.3919
8.2102 | -34.334 | -14.4018
8.2163 | -34.3344 | -14.4117
8.2224 | -34.3348 | -14.4217
8.2285 | -34.3351 | -14.4314
8.2346 | -34.3355 | -14.4414
8.2408 | -34.3358 | -14.4512
8.2469 | -34.3362 | -14.4612
8.253 | -34.3366 | -14.4711
8.2591 | -34.3369 | -14.4812
8.2652 | -34.3373 | -14.4912
8.2713 | -34.3377 | -14.5013
8.2774 | -34.3381 | -14.5113
8.2835 | -34.3384 | -14.5215
8.2896 | -34.3388 | -14.5313
8.2957 | -34.3392 | -14.5414
8.3018 | -34.3395 | -14.5514
8.3079 | -34.3399 | -14.5612
8.314 | -34.3403 | -14.5712
8.3201 | -34.3407 | -14.5811
8.3262 | -34.341 | -14.5912
8.3323 | -34.3414 | -14.601
8.3384 | -34.3417 | -14.6111
8.3445 | -34.3421 | -14.621
8.3506 | -34.3425 | -14.631
8.3567 | -34.3428 | -14.641
8.3628 | -34.3432 | -14.6511
8.3689 | -34.3436 | -14.6609
8.375 | -34.3439 | -14.6709
8.3812 | -34.3443 | -14.6807
8.3873 | -34.3446 | -14.6907
8.3934 | -34.345 | -14.7007
8.3995 | -34.3453 | -14.7106
8.4056 | -34.3457 | -14.7205
8.4117 | -34.3461 | -14.7305
8.4178 | -34.3464 | -14.7405
8.4239 | -34.3468 | -14.7503
8.43 | -34.3471 | -14.7603
8.4361 | -34.3475 | -14.7703
8.4422 | -34.3479 | -14.7801
8.4483 | -34.3482 | -14.79
8.4544 | -34.3486 | -14.8
8.4605 | -34.349 | -14.8098
8.4666 | -34.3493 | -14.8199
8.4727 | -34.3497 | -14.8299
8.4788 | -34.35 | -14.8397
8.4849 | -34.3504 | -14.8497
8.491 | -34.3507 | -14.8596
8.4971 | -34.3511 | -14.8694
8.5032 | -34.3515 | -14.8793
8.5093 | -34.3519 | -14.889
8.5154 | -34.3522 | -14.899
8.5215 | -34.3526 | -14.9088
8.5277 | -34.353 | -14.9188
8.5338 | -34.3533 | -14.9286
8.5399 | -34.3537 | -14.9385
8.546 | -34.3541 | -14.9483
8.5521 | -34.3544 | -14.9582
8.5582 | -34.3548 | -14.9681
8.5643 | -34.3551 | -14.9778
8.5704 | -34.3555 | -14.9875
8.5765 | -34.3559 | -14.9976
8.5826 | -34.3563 | -15.0071
8.5887 | -34.3566 | -15.0169
8.5948 | -34.357 | -15.0268
8.6009 | -34.3574 | -15.0365
8.607 | -34.3578 | -15.0465
8.6131 | -34.3581 | -15.0564
8.6192 | -34.3585 | -15.0661
8.6253 | -34.3589 | -15.0759
8.6314 | -34.3592 | -15.0857
8.6375 | -34.3596 | -15.0956
8.6436 | -34.36 | -15.1055
8.6497 | -34.3603 | -15.1153
8.6558 | -34.3607 | -15.1251
8.6619 | -34.361 | -15.135
8.6681 | -34.3614 | -15.145
8.6742 | -34.3618 | -15.1548
8.6803 | -34.3622 | -15.1646
8.6864 | -34.3626 | -15.1742
8.6925 | -34.3629 | -15.1839
8.6986 | -34.3633 | -15.1938
8.7047 | -34.3636 | -15.2037
8.7108 | -34.364 | -15.2136
8.7169 | -34.3644 | -15.2235
8.723 | -34.3648 | -15.2332
8.7291 | -34.3651 | -15.2429
8.7352 | -34.3655 | -15.2528
8.7413 | -34.3658 | -15.2627
8.7474 | -34.3662 | -15.2724
8.7535 | -34.3666 | -15.2823
8.7596 | -34.367 | -15.2921
8.7657 | -34.3673 | -15.3021
8.7718 | -34.3677 | -15.312
8.7779 | -34.3681 | -15.3218
8.784 | -34.3684 | -15.3317
8.7901 | -34.3689 | -15.3415
8.7962 | -34.3692 | -15.3511
8.8023 | -34.3696 | -15.3609
8.8084 | -34.37 | -15.3707
8.8146 | -34.3703 | -15.3805
8.8207 | -34.3707 | -15.3902
8.8268 | -34.3711 | -15.3998
8.8329 | -34.3714 | -15.4096
8.839 | -34.3718 | -15.4194
8.8451 | -34.3722 | -15.4291
8.8512 | -34.3726 | -15.4388
8.8573 | -34.3729 | -15.4486
8.8634 | -34.3733 | -15.4582
8.8695 | -34.3737 | -15.4682
8.8756 | -34.374 | -15.478
8.8817 | -34.3744 | -15.4877
8.8878 | -34.3748 | -15.4977
8.8939 | -34.3752 | -15.5075
8.9 | -34.3756 | -15.5172
8.9061 | -34.3759 | -15.5269
8.9122 | -34.3763 | -15.5367
8.9183 | -34.3767 | -15.5465
8.9244 | -34.3771 | -15.5565
8.9305 | -34.3775 | -15.566
8.9366 | -34.3779 | -15.5758
8.9427 | -34.3782 | -15.5856
8.9488 | -34.3786 | -15.5953
8.955 | -34.379 | -15.6051
8.9611 | -34.3794 | -15.615
8.9672 | -34.3798 | -15.6247
8.9733 | -34.3802 | -15.6343
8.9794 | -34.3806 | -15.6439
8.9855 | -34.381 | -15.6538
8.9916 | -34.3813 | -15.6637
8.9977 | -34.3817 | -15.6734
9.0038 | -34.3822 | -15.683
9.0099 | -34.3825 | -15.6927
9.016 | -34.3829 | -15.7025
9.0221 | -34.3833 | -15.7121
9.0282 | -34.3837 | -15.7216
9.0343 | -34.3841 | -15.7314
9.0404 | -34.3845 | -15.7411
9.0465 | -34.3849 | -15.751
9.0526 | -34.3853 | -15.7606
9.0587 | -34.3857 | -15.7704
9.0648 | -34.3861 | -15.78
9.0709 | -34.3865 | -15.7897
9.077 | -34.3869 | -15.7995
9.0831 | -34.3872 | -15.8093
9.0892 | -34.3876 | -15.8191
9.0953 | -34.388 | -15.8289
9.1015 | -34.3884 | -15.8386
9.1076 | -34.3888 | -15.8482
9.1137 | -34.3892 | -15.8579
9.1198 | -34.3896 | -15.8675
9.1259 | -34.39 | -15.8772
9.132 | -34.3904 | -15.8871
9.1381 | -34.3908 | -15.8967
9.1442 | -34.3913 | -15.9062
9.1503 | -34.3916 | -15.916
9.1564 | -34.392 | -15.9257
9.1625 | -34.3924 | -15.9355
9.1686 | -34.3928 | -15.9451
9.1747 | -34.3932 | -15.9548
9.1808 | -34.3937 | -15.9645
9.1869 | -34.3941 | -15.9742
9.193 | -34.3944 | -15.9839
9.1991 | -34.3948 | -15.9937
9.2052 | -34.3952 | -16.0033
9.2113 | -34.3956 | -16.0131
9.2174 | -34.3961 | -16.0227
9.2235 | -34.3965 | -16.0324
9.2296 | -34.3969 | -16.042
9.2357 | -34.3973 | -16.0519
9.2419 | -34.3977 | -16.0615
9.248 | -34.3981 | -16.0712
9.2541 | -34.3985 | -16.0809
9.2602 | -34.3989 | -16.0908
9.2663 | -34.3994 | -16.1002
9.2724 | -34.3998 | -16.1099
9.2785 | -34.4001 | -16.1196
9.2846 | -34.4006 | -16.1291
9.2907 | -34.401 | -16.1388
9.2968 | -34.4014 | -16.1486
9.3029 | -34.4018 | -16.1584
9.309 | -34.4022 | -16.1681
9.3151 | -34.4026 | -16.1777
9.3212 | -34.403 | -16.1872
9.3273 | -34.4034 | -16.1969
9.3334 | -34.4038 | -16.2065
9.3395 | -34.4042 | -16.216
9.3456 | -34.4046 | -16.2255
9.3517 | -34.405 | -16.2353
9.3578 | -34.4055 | -16.2449
9.3639 | -34.4059 | -16.2544
9.37 | -34.4063 | -16.2641
9.3761 | -34.4067 | -16.274
9.3822 | -34.4071 | -16.2837
9.3884 | -34.4076 | -16.2931
9.3945 | -34.408 | -16.3028
9.4006 | -34.4084 | -16.3125
9.4067 | -34.4088 | -16.3222
9.4128 | -34.4092 | -16.3319
9.4189 | -34.4096 | -16.3417
9.425 | -34.4101 | -16.3512
9.4311 | -34.4105 | -16.3609
9.4372 | -34.4109 | -16.3705
9.4433 | -34.4113 | -16.3801
9.4494 | -34.4117 | -16.3896
9.4555 | -34.4122 | -16.3994
9.4616 | -34.4126 | -16.409
9.4677 | -34.413 | -16.4186
9.4738 | -34.4134 | -16.4282
9.4799 | -34.4139 | -16.4379
9.486 | -34.4143 | -16.4476
9.4921 | -34.4147 | -16.4572
9.4982 | -34.4151 | -16.4666
9.5043 | -34.4155 | -16.4763
9.5104 | -34.4159 | -16.4858
9.5165 | -34.4164 | -16.4954
9.5226 | -34.4168 | -16.5052
9.5288 | -34.4172 | -16.5147
9.5349 | -34.4176 | -16.5244
9.541 | -34.418 | -16.5339
9.5471 | -34.4184 | -16.5435
9.5532 | -34.4189 | -16.5532
9.5593 | -34.4193 | -16.5629
9.5654 | -34.4197 | -16.5725
9.5715 | -34.4201 | -16.5821
9.5776 | -34.4206 | -16.5919
9.5837 | -34.4209 | -16.6014
9.5898 | -34.4213 | -16.6112
9.5959 | -34.4218 | -16.6209
9.602 | -34.4222 | -16.6303
9.6081 | -34.4226 | -16.6399
9.6142 | -34.423 | -16.6494
9.6203 | -34.4234 | -16.6591
9.6264 | -34.4238 | -16.6686
9.6325 | -34.4242 | -16.6783
9.6386 | -34.4247 | -16.6878
9.6447 | -34.4251 | -16.6975
9.6508 | -34.4255 | -16.707
9.6569 | -34.4259 | -16.7166
9.663 | -34.4264 | -16.7263
9.6691 | -34.4268 | -16.7359
9.6753 | -34.4273 | -16.7454
9.6814 | -34.4277 | -16.7549
9.6875 | -34.4281 | -16.7643
9.6936 | -34.4285 | -16.7738
9.6997 | -34.4289 | -16.7835
9.7058 | -34.4294 | -16.793
9.7119 | -34.4298 | -16.8027
9.718 | -34.4303 | -16.8124
9.7241 | -34.4307 | -16.8219
9.7302 | -34.4311 | -16.8315
9.7363 | -34.4315 | -16.841
9.7424 | -34.432 | -16.8504
9.7485 | -34.4324 | -16.8601
9.7546 | -34.4328 | -16.8695
9.7607 | -34.4332 | -16.8793
9.7668 | -34.4336 | -16.8889
9.7729 | -34.4341 | -16.8984
9.779 | -34.4345 | -16.9078
9.7851 | -34.4349 | -16.9175
9.7912 | -34.4354 | -16.927
9.7973 | -34.4358 | -16.9367
9.8034 | -34.4363 | -16.9461
9.8095 | -34.4367 | -16.9559
9.8157 | -34.4371 | -16.9654
9.8218 | -34.4375 | -16.9752
9.8279 | -34.4379 | -16.9846
9.834 | -34.4383 | -16.9943
9.8401 | -34.4388 | -17.0038
9.8462 | -34.4392 | -17.0132
9.8523 | -34.4397 | -17.0227
9.8584 | -34.4401 | -17.0323
9.8645 | -34.4405 | -17.042
9.8706 | -34.4409 | -17.0517
9.8767 | -34.4413 | -17.0612
9.8828 | -34.4417 | -17.0706
9.8889 | -34.4422 | -17.08
9.895 | -34.4426 | -17.0896
9.9011 | -34.443 | -17.0992
9.9072 | -34.4434 | -17.1088
9.9133 | -34.4438 | -17.1183
9.9194 | -34.4443 | -17.1277
9.9255 | -34.4447 | -17.1372
9.9316 | -34.4451 | -17.1468
9.9377 | -34.4455 | -17.1563
9.9438 | -34.4459 | -17.1659
9.9499 | -34.4464 | -17.1756
9.956 | -34.4468 | -17.1851
9.9622 | -34.4472 | -17.1948
9.9683 | -34.4477 | -17.2043
9.9744 | -34.4481 | -17.214
9.9805 | -34.4485 | -17.2235
9.9866 | -34.4489 | -17.233
9.9927 | -34.4493 | -17.2424
9.9988 | -34.4498 | -17.2521
10.0049 | -34.4502 | -17.2616
10.011 | -34.4506 | -17.2712
10.0171 | -34.4511 | -17.2806
10.0232 | -34.4515 | -17.29
10.0293 | -34.4519 | -17.2996
10.0354 | -34.4524 | -17.3092
10.0415 | -34.4528 | -17.3187
10.0476 | -34.4532 | -17.3284
10.0537 | -34.4537 | -17.338
10.0598 | -34.4541 | -17.3474
10.0659 | -34.4545 | -17.357
10.072 | -34.4549 | -17.3664
10.0781 | -34.4553 | -17.3758
10.0842 | -34.4558 | -17.3854
10.0903 | -34.4562 | -17.3949
10.0964 | -34.4566 | -17.4041
10.1026 | -34.457 | -17.4137
10.1087 | -34.4574 | -17.4232
10.1148 | -34.4578 | -17.4328
10.1209 | -34.4583 | -17.4422
10.127 | -34.4587 | -17.4516
10.1331 | -34.4591 | -17.461
10.1392 | -34.4596 | -17.4703
10.1453 | -34.46 | -17.4797
10.1514 | -34.4604 | -17.4893
10.1575 | -34.4608 | -17.4989
10.1636 | -34.4613 | -17.5082
10.1697 | -34.4617 | -17.5178
10.1758 | -34.4621 | -17.5271
10.1819 | -34.4626 | -17.5365
10.188 | -34.463 | -17.5461
10.1941 | -34.4634 | -17.5556
10.2002 | -34.4638 | -17.5649
10.2063 | -34.4642 | -17.5744
10.2124 | -34.4646 | -17.5839
10.2185 | -34.4651 | -17.5934
10.2246 | -34.4655 | -17.6028
10.2307 | -34.4659 | -17.6123
10.2368 | -34.4664 | -17.6215
10.2429 | -34.4668 | -17.6309
10.2491 | -34.4672 | -17.6405
10.2552 | -34.4676 | -17.6499
10.2613 | -34.468 | -17.6595
10.2674 | -34.4685 | -17.6689
10.2735 | -34.4689 | -17.6786
10.2796 | -34.4693 | -17.688
10.2857 | -34.4698 | -17.6975
10.2918 | -34.4702 | -17.7071
10.2979 | -34.4706 | -17.7166
10.304 | -34.4711 | -17.7258
10.3101 | -34.4715 | -17.7354
10.3162 | -34.4719 | -17.7449
10.3223 | -34.4723 | -17.7543
10.3284 | -34.4727 | -17.7637
10.3345 | -34.4732 | -17.7731
10.3406 | -34.4736 | -17.7825
10.3467 | -34.474 | -17.7918
10.3528 | -34.4744 | -17.8012
10.3589 | -34.4748 | -17.8107
10.365 | -34.4753 | -17.8202
10.3711 | -34.4757 | -17.8298
10.3772 | -34.4761 | -17.8392
10.3833 | -34.4765 | -17.8488
10.3895 | -34.477 | -17.8582
10.3956 | -34.4774 | -17.8677
10.4017 | -34.4779 | -17.877
10.4078 | -34.4783 | -17.8864
10.4139 | -34.4787 | -17.8959
10.42 | -34.4791 | -17.9054
10.4261 | -34.4796 | -17.9148
10.4322 | -34.48 | -17.9242
10.4383 | -34.4804 | -17.9336
10.4444 | -34.4808 | -17.9428
10.4505 | -34.4812 | -17.9523
10.4566 | -34.4816 | -17.9616
10.4627 | -34.482 | -17.971
10.4688 | -34.4825 | -17.9803
10.4749 | -34.4829 | -17.9898
10.481 | -34.4833 | -17.9992
10.4871 | -34.4838 | -18.0085
10.4932 | -34.4842 | -18.0178
10.4993 | -34.4846 | -18.0272
10.5054 | -34.485 | -18.0367
10.5115 | -34.4855 | -18.046
10.5176 | -34.4859 | -18.0554
10.5237 | -34.4864 | -18.0649
10.5298 | -34.4868 | -18.0742
10.536 | -34.4872 | -18.0837
10.5421 | -34.4876 | -18.093
10.5482 | -34.4881 | -18.1024
10.5543 | -34.4885 | -18.1116
10.5604 | -34.4889 | -18.1209
10.5665 | -34.4894 | -18.1302
10.5726 | -34.4898 | -18.1396
10.5787 | -34.4902 | -18.1489
10.5848 | -34.4906 | -18.158
10.5909 | -34.4911 | -18.1675
10.597 | -34.4915 | -18.1769
10.6031 | -34.4919 | -18.1864
10.6092 | -34.4923 | -18.1958
10.6153 | -34.4927 | -18.2052
10.6214 | -34.4932 | -18.2144
10.6275 | -34.4936 | -18.2238
10.6336 | -34.494 | -18.2332
10.6397 | -34.4944 | -18.2426
10.6458 | -34.4949 | -18.252
10.6519 | -34.4953 | -18.2614
10.658 | -34.4957 | -18.2706
10.6641 | -34.4961 | -18.2801
10.6702 | -34.4966 | -18.2894
10.6764 | -34.497 | -18.2988
10.6825 | -34.4974 | -18.3081
10.6886 | -34.4978 | -18.3173
10.6947 | -34.4983 | -18.3266
10.7008 | -34.4987 | -18.336
10.7069 | -34.4992 | -18.3455
10.713 | -34.4996 | -18.3548
10.7191 | -34.5 | -18.3641
10.7252 | -34.5005 | -18.3732
10.7313 | -34.5009 | -18.3824
10.7374 | -34.5013 | -18.3918
10.7435 | -34.5018 | -18.4011
10.7496 | -34.5022 | -18.4103
10.7557 | -34.5026 | -18.4197
10.7618 | -34.5031 | -18.4288
10.7679 | -34.5035 | -18.4381
10.774 | -34.5039 | -18.4475
10.7801 | -34.5043 | -18.4568
10.7862 | -34.5048 | -18.4661
10.7923 | -34.5052 | -18.4755
10.7984 | -34.5056 | -18.4846
10.8045 | -34.5061 | -18.4937
10.8106 | -34.5065 | -18.5031
10.8168 | -34.5069 | -18.5124
10.8229 | -34.5073 | -18.5217
10.829 | -34.5078 | -18.5309
10.8351 | -34.5082 | -18.5399
10.8412 | -34.5087 | -18.549
10.8473 | -34.5091 | -18.5582
10.8534 | -34.5095 | -18.5675
10.8595 | -34.51 | -18.5768
10.8656 | -34.5104 | -18.5859
10.8717 | -34.5108 | -18.5952
10.8778 | -34.5113 | -18.6044
10.8839 | -34.5117 | -18.6137
10.89 | -34.5121 | -18.6231
10.8961 | -34.5126 | -18.6323
10.9022 | -34.513 | -18.6415
10.9083 | -34.5134 | -18.6508
10.9144 | -34.5139 | -18.66
10.9205 | -34.5143 | -18.6692
10.9266 | -34.5147 | -18.6785
10.9327 | -34.5152 | -18.6877
10.9388 | -34.5156 | -18.6968
10.9449 | -34.516 | -18.706
10.951 | -34.5165 | -18.7153
10.9571 | -34.5169 | -18.7245
10.9633 | -34.5173 | -18.7338
10.9694 | -34.5177 | -18.7431
10.9755 | -34.5182 | -18.7523
10.9816 | -34.5186 | -18.7617
10.9877 | -34.5191 | -18.7709
10.9938 | -34.5195 | -18.7802
10.9999 | -34.52 | -18.7894
11.006 | -34.5204 | -18.7982
11.0121 | -34.5208 | -18.8074
11.0182 | -34.5213 | -18.8165
11.0243 | -34.5217 | -18.8257
11.0304 | -34.5221 | -18.835
11.0365 | -34.5226 | -18.8441
11.0426 | -34.5231 | -18.8533
11.0487 | -34.5235 | -18.8625
11.0548 | -34.524 | -18.8717
11.0609 | -34.5244 | -18.8809
11.067 | -34.5248 | -18.89
11.0731 | -34.5253 | -18.8991
11.0792 | -34.5257 | -18.9082
11.0853 | -34.5262 | -18.9173
11.0914 | -34.5266 | -18.9266
11.0975 | -34.527 | -18.9357
11.1037 | -34.5274 | -18.9448
11.1098 | -34.5279 | -18.9542
11.1159 | -34.5283 | -18.9631
11.122 | -34.5287 | -18.9723
11.1281 | -34.5292 | -18.9814
11.1342 | -34.5296 | -18.9906
11.1403 | -34.5301 | -18.9999
11.1464 | -34.5305 | -19.0092
11.1525 | -34.5309 | -19.0181
11.1586 | -34.5314 | -19.0274
11.1647 | -34.5318 | -19.0365
11.1708 | -34.5323 | -19.0457
11.1769 | -34.5327 | -19.0549
11.183 | -34.5332 | -19.0639
11.1891 | -34.5336 | -19.0731
11.1952 | -34.534 | -19.0823
11.2013 | -34.5345 | -19.0914
11.2074 | -34.5349 | -19.1005
11.2135 | -34.5354 | -19.11
11.2196 | -34.5358 | -19.1191
11.2257 | -34.5362 | -19.1282
11.2318 | -34.5367 | -19.1373
11.2379 | -34.5371 | -19.1465
11.244 | -34.5376 | -19.1557
11.2502 | -34.538 | -19.1649
11.2563 | -34.5385 | -19.1741
11.2624 | -34.5389 | -19.1833
11.2685 | -34.5394 | -19.1924
11.2746 | -34.5398 | -19.2015
11.2807 | -34.5403 | -19.2108
11.2868 | -34.5407 | -19.2201
11.2929 | -34.5412 | -19.229
11.299 | -34.5416 | -19.2382
11.3051 | -34.5421 | -19.2473
11.3112 | -34.5425 | -19.2565
11.3173 | -34.543 | -19.2654
11.3234 | -34.5434 | -19.2742
11.3295 | -34.5439 | -19.2834
11.3356 | -34.5443 | -19.2925
11.3417 | -34.5448 | -19.3015
11.3478 | -34.5452 | -19.3107
11.3539 | -34.5456 | -19.3199
11.36 | -34.546 | -19.3289
11.3661 | -34.5465 | -19.3378
11.3722 | -34.5469 | -19.3468
11.3783 | -34.5473 | -19.3562
11.3844 | -34.5478 | -19.3654
11.3906 | -34.5482 | -19.3746
11.3967 | -34.5487 | -19.3836
11.4028 | -34.5491 | -19.3928
11.4089 | -34.5496 | -19.4019
11.415 | -34.55 | -19.411
11.4211 | -34.5504 | -19.4201
11.4272 | -34.5508 | -19.4293
11.4333 | -34.5513 | -19.4384
11.4394 | -34.5517 | -19.4475
11.4455 | -34.5521 | -19.4565
11.4516 | -34.5526 | -19.4657
11.4577 | -34.553 | -19.475
11.4638 | -34.5535 | -19.4838
11.4699 | -34.5539 | -19.493
11.476 | -34.5544 | -19.5021
11.4821 | -34.5548 | -19.5112
11.4882 | -34.5553 | -19.5205
11.4943 | -34.5557 | -19.5298
11.5004 | -34.5562 | -19.5388
11.5065 | -34.5566 | -19.5477
11.5126 | -34.5571 | -19.5568
11.5187 | -34.5575 | -19.566
11.5248 | -34.5579 | -19.5754
11.5309 | -34.5584 | -19.5843
11.5371 | -34.5588 | -19.5933
11.5432 | -34.5593 | -19.6024
11.5493 | -34.5597 | -19.6116
11.5554 | -34.5602 | -19.6205
11.5615 | -34.5606 | -19.6298
11.5676 | -34.5611 | -19.6387
11.5737 | -34.5615 | -19.648
11.5798 | -34.5619 | -19.657
11.5859 | -34.5624 | -19.6662
11.592 | -34.5628 | -19.6751
11.5981 | -34.5632 | -19.6841
11.6042 | -34.5637 | -19.693
11.6103 | -34.5641 | -19.7021
11.6164 | -34.5646 | -19.7112
11.6225 | -34.565 | -19.7202
11.6286 | -34.5655 | -19.7292
11.6347 | -34.5659 | -19.7385
11.6408 | -34.5664 | -19.7474
11.6469 | -34.5668 | -19.7565
11.653 | -34.5672 | -19.7654
11.6591 | -34.5677 | -19.7744
11.6652 | -34.5681 | -19.7833
11.6713 | -34.5686 | -19.7923
11.6775 | -34.569 | -19.8013
11.6836 | -34.5695 | -19.8103
11.6897 | -34.57 | -19.8195
11.6958 | -34.5704 | -19.8284
11.7019 | -34.5709 | -19.8377
11.708 | -34.5713 | -19.8467
11.7141 | -34.5717 | -19.8556
11.7202 | -34.5722 | -19.8645
11.7263 | -34.5726 | -19.8736
11.7324 | -34.573 | -19.8828
11.7385 | -34.5735 | -19.8919
11.7446 | -34.5739 | -19.9008
11.7507 | -34.5743 | -19.9095
11.7568 | -34.5748 | -19.9184
11.7629 | -34.5752 | -19.9275
11.769 | -34.5756 | -19.9365
11.7751 | -34.5761 | -19.9455
11.7812 | -34.5765 | -19.9544
11.7873 | -34.577 | -19.9634
11.7934 | -34.5774 | -19.9724
11.7995 | -34.5779 | -19.9814
11.8056 | -34.5783 | -19.9902
11.8117 | -34.5788 | -19.9992
11.8178 | -34.5792 | -20.008
11.824 | -34.5797 | -20.017
11.8301 | -34.5801 | -20.0259
11.8362 | -34.5806 | -20.0347
11.8423 | -34.581 | -20.0438
11.8484 | -34.5814 | -20.053
11.8545 | -34.5819 | -20.0618
11.8606 | -34.5823 | -20.0708
11.8667 | -34.5828 | -20.08
11.8728 | -34.5832 | -20.0887
11.8789 | -34.5836 | -20.0977
11.885 | -34.5841 | -20.1066
11.8911 | -34.5845 | -20.1154
11.8972 | -34.585 | -20.1242
11.9033 | -34.5854 | -20.133
11.9094 | -34.5858 | -20.142
11.9155 | -34.5862 | -20.1511
11.9216 | -34.5867 | -20.1598
11.9277 | -34.5871 | -20.1687
11.9338 | -34.5875 | -20.1778
11.9399 | -34.588 | -20.187
11.946 | -34.5885 | -20.1957
11.9521 | -34.5889 | -20.2048
11.9582 | -34.5893 | -20.2134
11.9644 | -34.5897 | -20.2222
11.9705 | -34.5902 | -20.2312
11.9766 | -34.5906 | -20.2401
11.9827 | -34.5911 | -20.2491
11.9888 | -34.5916 | -20.258
11.9949 | -34.592 | -20.2668
12.001 | -34.5924 | -20.2758
12.0071 | -34.5929 | -20.2846
12.0132 | -34.5933 | -20.2934
12.0193 | -34.5937 | -20.3025
12.0254 | -34.5942 | -20.3113
12.0315 | -34.5947 | -20.3201
12.0376 | -34.5951 | -20.3289
12.0437 | -34.5955 | -20.3377
12.0498 | -34.596 | -20.3466
12.0559 | -34.5964 | -20.3555
12.062 | -34.5968 | -20.3643
12.0681 | -34.5973 | -20.3733
12.0742 | -34.5977 | -20.3822
12.0803 | -34.5981 | -20.3912
12.0864 | -34.5986 | -20.4001
12.0925 | -34.599 | -20.409
12.0986 | -34.5995 | -20.4177
12.1047 | -34.5999 | -20.4265
12.1109 | -34.6004 | -20.4353
12.117 | -34.6008 | -20.4441
12.1231 | -34.6013 | -20.4528
12.1292 | -34.6017 | -20.4618
12.1353 | -34.6022 | -20.4706
12.1414 | -34.6026 | -20.4795
12.1475 | -34.6031 | -20.4883
12.1536 | -34.6035 | -20.4971
12.1597 | -34.6039 | -20.5061
12.1658 | -34.6044 | -20.5149
12.1719 | -34.6048 | -20.5236
12.178 | -34.6052 | -20.5325
12.1841 | -34.6057 | -20.5414
12.1902 | -34.6062 | -20.5503
12.1963 | -34.6066 | -20.5592
12.2024 | -34.6071 | -20.5679
12.2085 | -34.6075 | -20.5767
12.2146 | -34.6079 | -20.5856
12.2207 | -34.6084 | -20.5944
12.2268 | -34.6089 | -20.6034
12.2329 | -34.6093 | -20.6122
12.239 | -34.6098 | -20.6209
12.2451 | -34.6102 | -20.6299
12.2513 | -34.6106 | -20.6388
12.2574 | -34.6111 | -20.6477
12.2635 | -34.6115 | -20.6565
12.2696 | -34.612 | -20.6653
12.2757 | -34.6124 | -20.674
12.2818 | -34.6129 | -20.6828
12.2879 | -34.6133 | -20.6917
12.294 | -34.6138 | -20.7002
12.3001 | -34.6142 | -20.7091
12.3062 | -34.6147 | -20.7178
12.3123 | -34.6151 | -20.7264
12.3184 | -34.6156 | -20.7352
12.3245 | -34.616 | -20.7438
12.3306 | -34.6164 | -20.7525
12.3367 | -34.6169 | -20.7614
12.3428 | -34.6173 | -20.7702
12.3489 | -34.6178 | -20.7787
12.355 | -34.6182 | -20.7876
12.3611 | -34.6187 | -20.7966
12.3672 | -34.6191 | -20.8053
12.3733 | -34.6196 | -20.8141
12.3794 | -34.62 | -20.8227
12.3855 | -34.6204 | -20.8314
12.3916 | -34.6209 | -20.8401
12.3978 | -34.6214 | -20.8489
12.4039 | -34.6218 | -20.8577
12.41 | -34.6222 | -20.8665
12.4161 | -34.6227 | -20.8751
12.4222 | -34.6232 | -20.8839
12.4283 | -34.6236 | -20.8926
12.4344 | -34.624 | -20.9015
12.4405 | -34.6245 | -20.9103
12.4466 | -34.6249 | -20.9189
12.4527 | -34.6254 | -20.9274
12.4588 | -34.6258 | -20.9362
12.4649 | -34.6263 | -20.9451
12.471 | -34.6267 | -20.9538
12.4771 | -34.6271 | -20.9626
12.4832 | -34.6276 | -20.9712
12.4893 | -34.628 | -20.9798
12.4954 | -34.6285 | -20.9884
12.5015 | -34.6289 | -20.9971
12.5076 | -34.6294 | -21.006
12.5137 | -34.6298 | -21.0148
12.5198 | -34.6303 | -21.0232
12.5259 | -34.6307 | -21.032
12.532 | -34.6312 | -21.0407
12.5382 | -34.6316 | -21.0495
12.5443 | -34.6321 | -21.0583
12.5504 | -34.6325 | -21.0671
12.5565 | -34.633 | -21.0758
12.5626 | -34.6334 | -21.0846
12.5687 | -34.6338 | -21.0934
12.5748 | -34.6343 | -21.1022
12.5809 | -34.6347 | -21.1109
12.587 | -34.6352 | -21.1197
12.5931 | -34.6357 | -21.1283
12.5992 | -34.6361 | -21.137
12.6053 | -34.6365 | -21.1457
12.6114 | -34.637 | -21.1544
12.6175 | -34.6374 | -21.1631
12.6236 | -34.6379 | -21.1718
12.6297 | -34.6383 | -21.1805
12.6358 | -34.6388 | -21.1891
12.6419 | -34.6392 | -21.1978
12.648 | -34.6397 | -21.2066
12.6541 | -34.6401 | -21.2153
12.6602 | -34.6406 | -21.2237
12.6663 | -34.641 | -21.2325
12.6724 | -34.6415 | -21.2411
12.6785 | -34.6419 | -21.2497
12.6847 | -34.6424 | -21.2582
12.6908 | -34.6428 | -21.267
12.6969 | -34.6433 | -21.2756
12.703 | -34.6438 | -21.2843
12.7091 | -34.6442 | -21.2931
12.7152 | -34.6447 | -21.3017
12.7213 | -34.6451 | -21.3104
12.7274 | -34.6456 | -21.319
12.7335 | -34.646 | -21.3273
12.7396 | -34.6465 | -21.3359
12.7457 | -34.6469 | -21.3447
12.7518 | -34.6474 | -21.3533
12.7579 | -34.6478 | -21.3621
12.764 | -34.6482 | -21.3708
12.7701 | -34.6487 | -21.3794
12.7762 | -34.6492 | -21.388
12.7823 | -34.6496 | -21.3966
12.7884 | -34.6501 | -21.4053
12.7945 | -34.6505 | -21.414
12.8006 | -34.651 | -21.4226
12.8067 | -34.6514 | -21.4312
12.8128 | -34.6519 | -21.4398
12.8189 | -34.6524 | -21.4484
12.8251 | -34.6528 | -21.4571
12.8312 | -34.6533 | -21.4659
12.8373 | -34.6537 | -21.4744
12.8434 | -34.6542 | -21.4831
12.8495 | -34.6546 | -21.4917
12.8556 | -34.6551 | -21.5003
12.8617 | -34.6555 | -21.509
12.8678 | -34.656 | -21.5178
12.8739 | -34.6564 | -21.5263
12.88 | -34.6569 | -21.5348
12.8861 | -34.6573 | -21.5434
12.8922 | -34.6578 | -21.5521
12.8983 | -34.6582 | -21.5607
12.9044 | -34.6587 | -21.5692
12.9105 | -34.6592 | -21.5777
12.9166 | -34.6596 | -21.5863
12.9227 | -34.6601 | -21.5949
12.9288 | -34.6605 | -21.6035
12.9349 | -34.661 | -21.6121
12.941 | -34.6614 | -21.6207
12.9471 | -34.6619 | -21.6294
12.9532 | -34.6624 | -21.6379
12.9593 | -34.6628 | -21.6467
12.9654 | -34.6633 | -21.6554
12.9716 | -34.6638 | -21.6641
12.9777 | -34.6642 | -21.6726
12.9838 | -34.6647 | -21.6811
12.9899 | -34.6652 | -21.6898
12.996 | -34.6656 | -21.6984
13.0021 | -34.6661 | -21.707
13.0082 | -34.6665 | -21.7157
13.0143 | -34.667 | -21.7241
13.0204 | -34.6675 | -21.7326
13.0265 | -34.6679 | -21.7411
13.0326 | -34.6684 | -21.7497
13.0387 | -34.6688 | -21.7584
13.0448 | -34.6693 | -21.767
13.0509 | -34.6697 | -21.7755
13.057 | -34.6702 | -21.784
13.0631 | -34.6707 | -21.7925
13.0692 | -34.6712 | -21.801
13.0753 | -34.6716 | -21.8097
13.0814 | -34.6721 | -21.8184
13.0875 | -34.6726 | -21.8268
13.0936 | -34.673 | -21.8355
13.0997 | -34.6735 | -21.844
13.1058 | -34.6739 | -21.8525
13.112 | -34.6744 | -21.8611
13.1181 | -34.6749 | -21.8696
13.1242 | -34.6753 | -21.8782
13.1303 | -34.6758 | -21.8868
13.1364 | -34.6763 | -21.8953
13.1425 | -34.6767 | -21.9039
13.1486 | -34.6772 | -21.9125
13.1547 | -34.6776 | -21.9209
13.1608 | -34.6781 | -21.9294
13.1669 | -34.6786 | -21.9379
13.173 | -34.679 | -21.9465
13.1791 | -34.6795 | -21.9549
13.1852 | -34.6799 | -21.9635
13.1913 | -34.6804 | -21.9719
13.1974 | -34.6809 | -21.9805
13.2035 | -34.6813 | -21.989
13.2096 | -34.6818 | -21.9975
13.2157 | -34.6823 | -22.006
13.2218 | -34.6827 | -22.0145
13.2279 | -34.6832 | -22.0229
13.234 | -34.6837 | -22.0314
13.2401 | -34.6841 | -22.04
13.2462 | -34.6846 | -22.0485
13.2524 | -34.6851 | -22.057
13.2585 | -34.6855 | -22.0654
13.2646 | -34.686 | -22.074
13.2707 | -34.6865 | -22.0825
13.2768 | -34.687 | -22.0912
13.2829 | -34.6874 | -22.0997
13.289 | -34.6879 | -22.1081
13.2951 | -34.6884 | -22.1165
13.3012 | -34.6888 | -22.1251
13.3073 | -34.6893 | -22.1337
13.3134 | -34.6898 | -22.1424
13.3195 | -34.6902 | -22.1509
13.3256 | -34.6907 | -22.1593
13.3317 | -34.6912 | -22.1676
13.3378 | -34.6916 | -22.1762
13.3439 | -34.6921 | -22.1848
13.35 | -34.6926 | -22.1934
13.3561 | -34.6931 | -22.2018
13.3622 | -34.6935 | -22.2102
13.3683 | -34.694 | -22.2187
13.3744 | -34.6945 | -22.227
13.3805 | -34.6949 | -22.2356
13.3866 | -34.6954 | -22.244
13.3927 | -34.6958 | -22.2526
13.3989 | -34.6963 | -22.2609
13.405 | -34.6968 | -22.2694
13.4111 | -34.6973 | -22.2778
13.4172 | -34.6977 | -22.2863
13.4233 | -34.6982 | -22.2949
13.4294 | -34.6987 | -22.3035
13.4355 | -34.6991 | -22.3118
13.4416 | -34.6996 | -22.3203
13.4477 | -34.7001 | -22.3287
13.4538 | -34.7006 | -22.3373
13.4599 | -34.7011 | -22.3458
13.466 | -34.7015 | -22.3544
13.4721 | -34.702 | -22.3627
13.4782 | -34.7025 | -22.3712
13.4843 | -34.703 | -22.3798
13.4904 | -34.7034 | -22.3882
13.4965 | -34.7039 | -22.3968
13.5026 | -34.7044 | -22.4053
13.5087 | -34.7048 | -22.4137
13.5148 | -34.7053 | -22.4223
13.5209 | -34.7058 | -22.4308
13.527 | -34.7063 | -22.4393
13.5331 | -34.7067 | -22.4478
13.5393 | -34.7072 | -22.4563
13.5454 | -34.7077 | -22.4646
13.5515 | -34.7082 | -22.4732
13.5576 | -34.7087 | -22.4817
13.5637 | -34.7091 | -22.4903
13.5698 | -34.7096 | -22.4988
13.5759 | -34.7101 | -22.5069
13.582 | -34.7106 | -22.5154
13.5881 | -34.7111 | -22.5238
13.5942 | -34.7115 | -22.5324
13.6003 | -34.712 | -22.5408
13.6064 | -34.7125 | -22.5493
13.6125 | -34.713 | -22.5578
13.6186 | -34.7135 | -22.5661
13.6247 | -34.714 | -22.5746
13.6308 | -34.7145 | -22.5831
13.6369 | -34.7149 | -22.5916
13.643 | -34.7154 | -22.6002
13.6491 | -34.7159 | -22.6085
13.6552 | -34.7164 | -22.6169
13.6613 | -34.7169 | -22.6253
13.6674 | -34.7173 | -22.6338
13.6735 | -34.7178 | -22.6423
13.6796 | -34.7183 | -22.6508
13.6858 | -34.7188 | -22.659
13.6919 | -34.7193 | -22.6674
13.698 | -34.7197 | -22.6759
13.7041 | -34.7202 | -22.6844
13.7102 | -34.7207 | -22.6928
13.7163 | -34.7212 | -22.7012
13.7224 | -34.7217 | -22.7096
13.7285 | -34.7222 | -22.7181
13.7346 | -34.7226 | -22.7266
13.7407 | -34.7231 | -22.7349
13.7468 | -34.7236 | -22.7433
13.7529 | -34.7241 | -22.7516
13.759 | -34.7246 | -22.7602
13.7651 | -34.7251 | -22.7686
13.7712 | -34.7255 | -22.777
13.7773 | -34.726 | -22.7855
13.7834 | -34.7265 | -22.7939
13.7895 | -34.727 | -22.8024
13.7956 | -34.7275 | -22.8109
13.8017 | -34.728 | -22.8193
13.8078 | -34.7284 | -22.8278
13.8139 | -34.7289 | -22.8361
13.82 | -34.7294 | -22.8445
13.8262 | -34.7299 | -22.853
13.8323 | -34.7304 | -22.8615
13.8384 | -34.7309 | -22.87
13.8445 | -34.7314 | -22.8784
13.8506 | -34.7319 | -22.8868
13.8567 | -34.7323 | -22.8952
13.8628 | -34.7328 | -22.9037
13.8689 | -34.7333 | -22.9123
13.875 | -34.7338 | -22.9207
13.8811 | -34.7342 | -22.9292
13.8872 | -34.7347 | -22.9376
13.8933 | -34.7352 | -22.9459
13.8994 | -34.7357 | -22.9543
13.9055 | -34.7362 | -22.9629
13.9116 | -34.7367 | -22.9712
13.9177 | -34.7371 | -22.9797
13.9238 | -34.7376 | -22.9883
13.9299 | -34.7381 | -22.9966
13.936 | -34.7386 | -23.005
13.9421 | -34.7391 | -23.0134
13.9482 | -34.7395 | -23.0218
13.9543 | -34.74 | -23.0303
13.9604 | -34.7405 | -23.0387
13.9665 | -34.741 | -23.047
13.9727 | -34.7415 | -23.0556
13.9788 | -34.742 | -23.0639
13.9849 | -34.7425 | -23.0723
13.991 | -34.743 | -23.0807
13.9971 | -34.7435 | -23.0889
14.0032 | -34.7439 | -23.0973
14.0093 | -34.7444 | -23.1056
14.0154 | -34.7449 | -23.114
14.0215 | -34.7454 | -23.1223
14.0276 | -34.7459 | -23.1307
14.0337 | -34.7464 | -23.1391
14.0398 | -34.7469 | -23.1474
14.0459 | -34.7474 | -23.1558
14.052 | -34.7478 | -23.1643
14.0581 | -34.7483 | -23.1727
14.0642 | -34.7488 | -23.1811
14.0703 | -34.7493 | -23.1893
14.0764 | -34.7498 | -23.1976
14.0825 | -34.7503 | -23.206
14.0886 | -34.7508 | -23.2144
14.0947 | -34.7513 | -23.2228
14.1008 | -34.7518 | -23.2312
14.1069 | -34.7523 | -23.2394
14.1131 | -34.7527 | -23.2479
14.1192 | -34.7533 | -23.2562
14.1253 | -34.7537 | -23.2645
14.1314 | -34.7542 | -23.2728
14.1375 | -34.7547 | -23.2811
14.1436 | -34.7552 | -23.2895
14.1497 | -34.7557 | -23.2978
14.1558 | -34.7562 | -23.3062
14.1619 | -34.7567 | -23.3145
14.168 | -34.7572 | -23.3229
14.1741 | -34.7577 | -23.3311
14.1802 | -34.7582 | -23.3394
14.1863 | -34.7586 | -23.3479
14.1924 | -34.7591 | -23.3563
14.1985 | -34.7596 | -23.3646
14.2046 | -34.7601 | -23.3731
14.2107 | -34.7606 | -23.3813
14.2168 | -34.7611 | -23.3896
14.2229 | -34.7616 | -23.3981
14.229 | -34.7621 | -23.4065
14.2351 | -34.7626 | -23.4149
14.2412 | -34.7631 | -23.4233
14.2473 | -34.7636 | -23.4315
14.2534 | -34.764 | -23.4397
14.2596 | -34.7645 | -23.4479
14.2657 | -34.765 | -23.4563
14.2718 | -34.7655 | -23.4647
14.2779 | -34.766 | -23.4731
14.284 | -34.7665 | -23.4816
14.2901 | -34.767 | -23.49
14.2962 | -34.7675 | -23.4984
14.3023 | -34.768 | -23.5068
14.3084 | -34.7685 | -23.5152
14.3145 | -34.769 | -23.5234
14.3206 | -34.7695 | -23.5319
14.3267 | -34.77 | -23.5403
14.3328 | -34.7705 | -23.5486
14.3389 | -34.7709 | -23.5569
14.345 | -34.7714 | -23.5653
14.3511 | -34.7719 | -23.5736
14.3572 | -34.7724 | -23.582
14.3633 | -34.7729 | -23.5904
14.3694 | -34.7734 | -23.5986
14.3755 | -34.7739 | -23.6069
14.3816 | -34.7744 | -23.6153
14.3877 | -34.7749 | -23.6235
14.3938 | -34.7754 | -23.632
14.4 | -34.7759 | -23.6402
14.4061 | -34.7764 | -23.6486
14.4122 | -34.7769 | -23.657
14.4183 | -34.7774 | -23.6653
14.4244 | -34.7779 | -23.6735
14.4305 | -34.7784 | -23.6818
14.4366 | -34.7789 | -23.6901
14.4427 | -34.7794 | -23.6985
14.4488 | -34.7799 | -23.7068
14.4549 | -34.7804 | -23.715
14.461 | -34.7809 | -23.7232
14.4671 | -34.7813 | -23.7317
14.4732 | -34.7818 | -23.74
14.4793 | -34.7823 | -23.7484
14.4854 | -34.7828 | -23.7566
14.4915 | -34.7833 | -23.7647
14.4976 | -34.7838 | -23.773
14.5037 | -34.7843 | -23.7812
14.5098 | -34.7848 | -23.7895
14.5159 | -34.7853 | -23.7979
14.522 | -34.7858 | -23.8061
14.5281 | -34.7863 | -23.8143
14.5342 | -34.7868 | -23.8227
14.5403 | -34.7873 | -23.8309
14.5465 | -34.7878 | -23.8393
14.5526 | -34.7882 | -23.8477
14.5587 | -34.7887 | -23.8561
14.5648 | -34.7892 | -23.8642
14.5709 | -34.7897 | -23.8725
14.577 | -34.7902 | -23.8807
14.5831 | -34.7907 | -23.8891
14.5892 | -34.7912 | -23.8973
14.5953 | -34.7917 | -23.9054
14.6014 | -34.7922 | -23.9138
14.6075 | -34.7926 | -23.922
14.6136 | -34.7931 | -23.9303
14.6197 | -34.7936 | -23.9387
14.6258 | -34.7941 | -23.947
14.6319 | -34.7946 | -23.9551
14.638 | -34.7951 | -23.9634
14.6441 | -34.7955 | -23.9717
14.6502 | -34.796 | -23.9801
14.6563 | -34.7965 | -23.9883
14.6624 | -34.797 | -23.9966
14.6685 | -34.7975 | -24.0046
14.6746 | -34.798 | -24.013
14.6807 | -34.7985 | -24.0211
14.6869 | -34.799 | -24.0294
14.693 | -34.7995 | -24.0377
14.6991 | -34.8 | -24.0459
14.7052 | -34.8005 | -24.0541
14.7113 | -34.801 | -24.0624
14.7174 | -34.8015 | -24.0707
14.7235 | -34.802 | -24.0789
14.7296 | -34.8025 | -24.0871
14.7357 | -34.803 | -24.0953
14.7418 | -34.8035 | -24.1035
14.7479 | -34.804 | -24.1116
14.754 | -34.8045 | -24.1198
14.7601 | -34.805 | -24.1279
14.7662 | -34.8054 | -24.1361
14.7723 | -34.8059 | -24.144
14.7784 | -34.8064 | -24.1524
14.7845 | -34.8069 | -24.1606
14.7906 | -34.8074 | -24.169
14.7967 | -34.8079 | -24.1772
14.8028 | -34.8084 | -24.1853
14.8089 | -34.8089 | -24.1935
14.815 | -34.8094 | -24.2018
14.8211 | -34.8099 | -24.2101
14.8272 | -34.8104 | -24.2185
14.8334 | -34.8109 | -24.2266
14.8395 | -34.8114 | -24.2348
14.8456 | -34.8119 | -24.243
14.8517 | -34.8123 | -24.2512
14.8578 | -34.8128 | -24.2596
14.8639 | -34.8133 | -24.2679
14.87 | -34.8138 | -24.2761
14.8761 | -34.8143 | -24.2843
14.8822 | -34.8148 | -24.2926
14.8883 | -34.8153 | -24.3009
14.8944 | -34.8158 | -24.309
14.9005 | -34.8163 | -24.3173
14.9066 | -34.8168 | -24.3255
14.9127 | -34.8173 | -24.3336
14.9188 | -34.8178 | -24.3417
14.9249 | -34.8183 | -24.35
14.931 | -34.8188 | -24.3581
14.9371 | -34.8193 | -24.3663
14.9432 | -34.8197 | -24.3745
14.9493 | -34.8202 | -24.3826
14.9554 | -34.8207 | -24.3908
14.9615 | -34.8212 | -24.399
14.9676 | -34.8217 | -24.4072
14.9738 | -34.8222 | -24.4154
14.9799 | -34.8227 | -24.4238
14.986 | -34.8232 | -24.4319
14.9921 | -34.8236 | -24.4401
14.9982 | -34.8241 | -24.4483
15.0043 | -34.8246 | -24.4564
15.0104 | -34.8251 | -24.4647
15.0165 | -34.8256 | -24.473
15.0226 | -34.8261 | -24.4811
15.0287 | -34.8266 | -24.4893
15.0348 | -34.8271 | -24.4975
15.0409 | -34.8276 | -24.5058
15.047 | -34.8281 | -24.5141
15.0531 | -34.8286 | -24.5221
15.0592 | -34.8291 | -24.5304
15.0653 | -34.8296 | -24.5385
15.0714 | -34.8301 | -24.5467
15.0775 | -34.8306 | -24.5549
15.0836 | -34.8311 | -24.5631
15.0897 | -34.8316 | -24.5711
15.0958 | -34.8321 | -24.5792
15.1019 | -34.8326 | -24.5874
15.108 | -34.8331 | -24.5955
15.1141 | -34.8336 | -24.6037
15.1203 | -34.8341 | -24.6119
15.1264 | -34.8346 | -24.62
15.1325 | -34.8351 | -24.6281
15.1386 | -34.8356 | -24.6363
15.1447 | -34.8361 | -24.6444
15.1508 | -34.8366 | -24.6526
15.1569 | -34.8371 | -24.6608
15.163 | -34.8376 | -24.6688
15.1691 | -34.8381 | -24.6769
15.1752 | -34.8386 | -24.685
15.1813 | -34.8391 | -24.6931
15.1874 | -34.8396 | -24.7013
15.1935 | -34.8401 | -24.7093
15.1996 | -34.8406 | -24.7174
15.2057 | -34.841 | -24.7255
15.2118 | -34.8415 | -24.7337
15.2179 | -34.842 | -24.7418
15.224 | -34.8425 | -24.75
15.2301 | -34.843 | -24.7579
15.2362 | -34.8435 | -24.7661
15.2423 | -34.844 | -24.7743
15.2484 | -34.8445 | -24.7824
15.2545 | -34.845 | -24.7906
15.2607 | -34.8455 | -24.7987
15.2668 | -34.846 | -24.8066
15.2729 | -34.8465 | -24.8147
15.279 | -34.847 | -24.8229
15.2851 | -34.8475 | -24.831
15.2912 | -34.848 | -24.8393
15.2973 | -34.8485 | -24.8475
15.3034 | -34.849 | -24.8556
15.3095 | -34.8495 | -24.8637
15.3156 | -34.85 | -24.8718
15.3217 | -34.8505 | -24.8799
15.3278 | -34.851 | -24.8881
15.3339 | -34.8515 | -24.896
15.34 | -34.852 | -24.904
15.3461 | -34.8525 | -24.9121
15.3522 | -34.853 | -24.9202
15.3583 | -34.8535 | -24.9285
15.3644 | -34.854 | -24.9366
15.3705 | -34.8545 | -24.9444
15.3766 | -34.855 | -24.9525
15.3827 | -34.8554 | -24.9605
15.3888 | -34.8559 | -24.9686
15.3949 | -34.8564 | -24.9768
15.401 | -34.8569 | -24.985
15.4072 | -34.8574 | -24.9929
15.4133 | -34.8579 | -25.001
15.4194 | -34.8584 | -25.0091
15.4255 | -34.8589 | -25.0172
15.4316 | -34.8594 | -25.0253
15.4377 | -34.8599 | -25.0335
15.4438 | -34.8604 | -25.0416
15.4499 | -34.8609 | -25.0495
15.456 | -34.8614 | -25.0574
15.4621 | -34.8619 | -25.0655
15.4682 | -34.8624 | -25.0737
15.4743 | -34.8629 | -25.0815
15.4804 | -34.8633 | -25.0896
15.4865 | -34.8638 | -25.0976
15.4926 | -34.8643 | -25.1058
15.4987 | -34.8648 | -25.1138
15.5048 | -34.8654 | -25.1219
15.5109 | -34.8658 | -25.1298
15.517 | -34.8664 | -25.1379
15.5231 | -34.8669 | -25.146
15.5292 | -34.8673 | -25.1539
15.5353 | -34.8678 | -25.162
15.5414 | -34.8683 | -25.17
15.5476 | -34.8688 | -25.1779
15.5537 | -34.8693 | -25.186
15.5598 | -34.8698 | -25.1941
15.5659 | -34.8703 | -25.2021
15.572 | -34.8708 | -25.2101
15.5781 | -34.8713 | -25.2182
15.5842 | -34.8718 | -25.226
15.5903 | -34.8723 | -25.234
15.5964 | -34.8728 | -25.2421
15.6025 | -34.8733 | -25.25
15.6086 | -34.8738 | -25.258
15.6147 | -34.8743 | -25.2661
15.6208 | -34.8748 | -25.274
15.6269 | -34.8753 | -25.282
15.633 | -34.8758 | -25.29
15.6391 | -34.8763 | -25.2982
15.6452 | -34.8767 | -25.3064
15.6513 | -34.8773 | -25.3141
15.6574 | -34.8778 | -25.3222
15.6635 | -34.8783 | -25.3302
15.6696 | -34.8788 | -25.3382
15.6757 | -34.8792 | -25.3463
15.6818 | -34.8797 | -25.3543
15.688 | -34.8802 | -25.3622
15.6941 | -34.8807 | -25.3701
15.7002 | -34.8812 | -25.3781
15.7063 | -34.8817 | -25.386
15.7124 | -34.8822 | -25.394
15.7185 | -34.8827 | -25.402
15.7246 | -34.8832 | -25.41
15.7307 | -34.8837 | -25.418
15.7368 | -34.8842 | -25.426
15.7429 | -34.8847 | -25.434
15.749 | -34.8852 | -25.4422
15.7551 | -34.8857 | -25.4503
15.7612 | -34.8862 | -25.4582
15.7673 | -34.8867 | -25.4661
15.7734 | -34.8872 | -25.4741
15.7795 | -34.8877 | -25.4822
15.7856 | -34.8882 | -25.49
15.7917 | -34.8887 | -25.4978
15.7978 | -34.8892 | -25.506
15.8039 | -34.8897 | -25.5139
15.81 | -34.8902 | -25.5219
15.8161 | -34.8907 | -25.5298
15.8222 | -34.8912 | -25.5378
15.8283 | -34.8917 | -25.5456
15.8345 | -34.8922 | -25.5535
15.8406 | -34.8927 | -25.5614
15.8467 | -34.8932 | -25.5694
15.8528 | -34.8937 | -25.5774
15.8589 | -34.8942 | -25.5853
15.865 | -34.8947 | -25.5932
15.8711 | -34.8952 | -25.6011
15.8772 | -34.8957 | -25.6091
15.8833 | -34.8962 | -25.6171
15.8894 | -34.8967 | -25.6251
15.8955 | -34.8972 | -25.633
15.9016 | -34.8977 | -25.6408
15.9077 | -34.8982 | -25.6488
15.9138 | -34.8987 | -25.6567
15.9199 | -34.8992 | -25.6647
15.926 | -34.8997 | -25.6726
15.9321 | -34.9002 | -25.6805
15.9382 | -34.9007 | -25.6882
15.9443 | -34.9011 | -25.6961
15.9504 | -34.9017 | -25.7042
15.9565 | -34.9022 | -25.7123
15.9626 | -34.9026 | -25.7203
15.9687 | -34.9031 | -25.728
15.9749 | -34.9037 | -25.7358
15.981 | -34.9041 | -25.7439
15.9871 | -34.9046 | -25.7519
15.9932 | -34.9051 | -25.7598
15.9993 | -34.9056 | -25.7677
16.0054 | -34.9061 | -25.7752
16.0115 | -34.9066 | -25.7832
16.0176 | -34.9071 | -25.7911
16.0237 | -34.9076 | -25.7991
16.0298 | -34.9081 | -25.807
16.0359 | -34.9086 | -25.8149
16.042 | -34.9091 | -25.8227
16.0481 | -34.9096 | -25.8306
16.0542 | -34.9101 | -25.8386
16.0603 | -34.9105 | -25.8465
16.0664 | -34.911 | -25.8544
16.0725 | -34.9115 | -25.8621
16.0786 | -34.912 | -25.87
16.0847 | -34.9125 | -25.8781
16.0908 | -34.913 | -25.8861
16.0969 | -34.9135 | -25.8942
16.103 | -34.914 | -25.902
16.1091 | -34.9145 | -25.9098
16.1152 | -34.915 | -25.9177
16.1214 | -34.9155 | -25.9256
16.1275 | -34.916 | -25.9335
16.1336 | -34.9165 | -25.9414
16.1397 | -34.917 | -25.9492
16.1458 | -34.9175 | -25.957
16.1519 | -34.918 | -25.965
16.158 | -34.9185 | -25.9731
16.1641 | -34.919 | -25.981
16.1702 | -34.9195 | -25.9888
16.1763 | -34.92 | -25.9967
16.1824 | -34.9205 | -26.0043
16.1885 | -34.921 | -26.012
16.1946 | -34.9215 | -26.0198
16.2007 | -34.922 | -26.0277
16.2068 | -34.9225 | -26.0357
16.2129 | -34.923 | -26.0436
16.219 | -34.9235 | -26.0512
16.2251 | -34.924 | -26.0589
16.2312 | -34.9245 | -26.0667
16.2373 | -34.925 | -26.0744
16.2434 | -34.9255 | -26.0822
16.2495 | -34.926 | -26.09
16.2556 | -34.9265 | -26.0979
16.2618 | -34.927 | -26.1056
16.2679 | -34.9275 | -26.1135
16.274 | -34.928 | -26.1212
16.2801 | -34.9285 | -26.1288
16.2862 | -34.929 | -26.1365
16.2923 | -34.9295 | -26.1445
16.2984 | -34.93 | -26.1524
16.3045 | -34.9305 | -26.1602
16.3106 | -34.931 | -26.1681
16.3167 | -34.9315 | -26.1759
16.3228 | -34.9319 | -26.1836
16.3289 | -34.9324 | -26.1914
16.335 | -34.9329 | -26.199
16.3411 | -34.9334 | -26.2069
16.3472 | -34.9339 | -26.2146
16.3533 | -34.9344 | -26.2225
16.3594 | -34.9349 | -26.2301
16.3655 | -34.9354 | -26.2379
16.3716 | -34.9359 | -26.2456
16.3777 | -34.9364 | -26.2533
16.3838 | -34.9369 | -26.2611
16.3899 | -34.9374 | -26.2689
16.396 | -34.9379 | -26.2767
16.4021 | -34.9383 | -26.2845
16.4083 | -34.9388 | -26.2924
16.4144 | -34.9393 | -26.3003
16.4205 | -34.9398 | -26.3083
16.4266 | -34.9403 | -26.316
16.4327 | -34.9408 | -26.3237
16.4388 | -34.9413 | -26.3316
16.4449 | -34.9418 | -26.3392
16.451 | -34.9423 | -26.347
16.4571 | -34.9428 | -26.355
16.4632 | -34.9433 | -26.3627
16.4693 | -34.9438 | -26.3705
16.4754 | -34.9443 | -26.3785
16.4815 | -34.9448 | -26.3864
16.4876 | -34.9453 | -26.3941
16.4937 | -34.9458 | -26.402
16.4998 | -34.9463 | -26.4097
16.5059 | -34.9468 | -26.4174
16.512 | -34.9472 | -26.4251
16.5181 | -34.9478 | -26.433
16.5242 | -34.9482 | -26.4407
16.5303 | -34.9487 | -26.4483
16.5364 | -34.9492 | -26.4558
16.5425 | -34.9497 | -26.4637
16.5487 | -34.9502 | -26.4716
16.5548 | -34.9507 | -26.4794
16.5609 | -34.9512 | -26.4872
16.567 | -34.9517 | -26.4948
16.5731 | -34.9521 | -26.5028
16.5792 | -34.9526 | -26.5106
16.5853 | -34.9531 | -26.5183
16.5914 | -34.9536 | -26.5262
16.5975 | -34.9541 | -26.534
16.6036 | -34.9546 | -26.5415
16.6097 | -34.9551 | -26.5492
16.6158 | -34.9555 | -26.5569
16.6219 | -34.956 | -26.5649
16.628 | -34.9565 | -26.5725
16.6341 | -34.957 | -26.5802
16.6402 | -34.9575 | -26.5877
16.6463 | -34.958 | -26.5955
16.6524 | -34.9585 | -26.6033
16.6585 | -34.959 | -26.6112
16.6646 | -34.9595 | -26.619
16.6707 | -34.96 | -26.6265
16.6768 | -34.9605 | -26.6342
16.6829 | -34.9609 | -26.6419
16.689 | -34.9615 | -26.6497
16.6952 | -34.962 | -26.6575
16.7013 | -34.9625 | -26.6652
16.7074 | -34.9629 | -26.6728
16.7135 | -34.9635 | -26.6806
16.7196 | -34.9639 | -26.6884
16.7257 | -34.9645 | -26.6961
16.7318 | -34.9649 | -26.7038
16.7379 | -34.9654 | -26.7114
16.744 | -34.9659 | -26.7192
16.7501 | -34.9664 | -26.7267
16.7562 | -34.9669 | -26.7346
16.7623 | -34.9673 | -26.7422
16.7684 | -34.9678 | -26.75
16.7745 | -34.9683 | -26.7579
16.7806 | -34.9689 | -26.7652
16.7867 | -34.9694 | -26.773
16.7928 | -34.9698 | -26.7807
16.7989 | -34.9703 | -26.7885
16.805 | -34.9708 | -26.7963
16.8111 | -34.9713 | -26.8041
16.8172 | -34.9718 | -26.8118
16.8233 | -34.9723 | -26.8196
16.8294 | -34.9728 | -26.8274
16.8356 | -34.9733 | -26.835
16.8417 | -34.9738 | -26.8428
16.8478 | -34.9743 | -26.8504
16.8539 | -34.9748 | -26.8579
16.86 | -34.9753 | -26.8657
16.8661 | -34.9758 | -26.8732
16.8722 | -34.9763 | -26.8809
16.8783 | -34.9768 | -26.8887
16.8844 | -34.9773 | -26.8963
16.8905 | -34.9778 | -26.9041
16.8966 | -34.9783 | -26.9117
16.9027 | -34.9788 | -26.9195
16.9088 | -34.9793 | -26.9272
16.9149 | -34.9798 | -26.935
16.921 | -34.9803 | -26.9424
16.9271 | -34.9808 | -26.9503
16.9332 | -34.9813 | -26.9579
16.9393 | -34.9818 | -26.9657
16.9454 | -34.9822 | -26.9733
16.9515 | -34.9827 | -26.9812
16.9576 | -34.9833 | -26.9887
16.9637 | -34.9838 | -26.9963
16.9698 | -34.9843 | -27.0038
16.9759 | -34.9847 | -27.0115
16.9821 | -34.9853 | -27.0193
16.9882 | -34.9858 | -27.0266
16.9943 | -34.9863 | -27.0342
17.0004 | -34.9868 | -27.042
17.0065 | -34.9872 | -27.0497
17.0126 | -34.9877 | -27.0572
17.0187 | -34.9882 | -27.0651
17.0248 | -34.9888 | -27.0725
17.0309 | -34.9893 | -27.0802
17.037 | -34.9897 | -27.0877
17.0431 | -34.9902 | -27.0954
17.0492 | -34.9907 | -27.103
17.0553 | -34.9912 | -27.1106
17.0614 | -34.9917 | -27.1183
17.0675 | -34.9922 | -27.126
17.0736 | -34.9926 | -27.1337
17.0797 | -34.9931 | -27.1413
17.0858 | -34.9936 | -27.1491
17.0919 | -34.9941 | -27.1567
17.098 | -34.9946 | -27.1643
17.1041 | -34.9951 | -27.1718
17.1102 | -34.9956 | -27.1795
17.1163 | -34.9961 | -27.1872
17.1225 | -34.9966 | -27.1947
17.1286 | -34.9971 | -27.2022
17.1347 | -34.9976 | -27.2097
17.1408 | -34.9981 | -27.2174
17.1469 | -34.9986 | -27.2249
17.153 | -34.9991 | -27.2324
17.1591 | -34.9996 | -27.24
17.1652 | -35.0001 | -27.2474
17.1713 | -35.0006 | -27.2551
17.1774 | -35.001 | -27.2627
17.1835 | -35.0015 | -27.2704
17.1896 | -35.002 | -27.2778
17.1957 | -35.0025 | -27.2856
17.2018 | -35.003 | -27.2929
17.2079 | -35.0035 | -27.3006
17.214 | -35.004 | -27.3083
17.2201 | -35.0045 | -27.316
17.2262 | -35.005 | -27.3236
17.2323 | -35.0055 | -27.3312
17.2384 | -35.006 | -27.3387
17.2445 | -35.0065 | -27.3462
17.2506 | -35.007 | -27.3538
17.2567 | -35.0075 | -27.3615
17.2628 | -35.0079 | -27.3692
17.269 | -35.0084 | -27.3766
17.2751 | -35.0089 | -27.3843
17.2812 | -35.0094 | -27.3919
17.2873 | -35.0099 | -27.3997
17.2934 | -35.0104 | -27.4072
17.2995 | -35.0109 | -27.4149
17.3056 | -35.0114 | -27.4224
17.3117 | -35.0119 | -27.43
17.3178 | -35.0124 | -27.4375
17.3239 | -35.0128 | -27.445
17.33 | -35.0133 | -27.4527
17.3361 | -35.0138 | -27.4602
17.3422 | -35.0143 | -27.4677
17.3483 | -35.0148 | -27.4755
17.3544 | -35.0153 | -27.483
17.3605 | -35.0158 | -27.4906
17.3666 | -35.0163 | -27.4981
17.3727 | -35.0168 | -27.5057
17.3788 | -35.0172 | -27.513
17.3849 | -35.0177 | -27.5205
17.391 | -35.0182 | -27.5281
17.3971 | -35.0187 | -27.5356
17.4032 | -35.0192 | -27.5432
17.4094 | -35.0197 | -27.5505
17.4155 | -35.0202 | -27.5581
17.4216 | -35.0207 | -27.5656
17.4277 | -35.0212 | -27.5732
17.4338 | -35.0217 | -27.5808
17.4399 | -35.0222 | -27.5884
17.446 | -35.0227 | -27.5958
17.4521 | -35.0232 | -27.6033
17.4582 | -35.0237 | -27.6107
17.4643 | -35.0242 | -27.6182
17.4704 | -35.0246 | -27.6259
17.4765 | -35.0251 | -27.6336
17.4826 | -35.0256 | -27.6413
17.4887 | -35.0261 | -27.6491
17.4948 | -35.0266 | -27.6566
17.5009 | -35.0271 | -27.6643
17.507 | -35.0276 | -27.6719
17.5131 | -35.0281 | -27.6794
17.5192 | -35.0286 | -27.6869
17.5253 | -35.0291 | -27.6946
17.5314 | -35.0296 | -27.702
17.5375 | -35.03 | -27.7096
17.5436 | -35.0305 | -27.7172
17.5497 | -35.031 | -27.7248
17.5559 | -35.0315 | -27.7321
17.562 | -35.032 | -27.7395
17.5681 | -35.0325 | -27.747
17.5742 | -35.033 | -27.7546
17.5803 | -35.0335 | -27.7622
17.5864 | -35.0341 | -27.7696
17.5925 | -35.0346 | -27.7773
17.5986 | -35.0351 | -27.7849
17.6047 | -35.0356 | -27.7925
17.6108 | -35.036 | -27.7999
17.6169 | -35.0365 | -27.8075
17.623 | -35.037 | -27.815
17.6291 | -35.0375 | -27.8226
17.6352 | -35.038 | -27.8301
17.6413 | -35.0385 | -27.8376
17.6474 | -35.039 | -27.8451
17.6535 | -35.0395 | -27.8524
17.6596 | -35.04 | -27.8598
17.6657 | -35.0405 | -27.8674
17.6718 | -35.0409 | -27.875
17.6779 | -35.0414 | -27.8824
17.684 | -35.0419 | -27.89
17.6901 | -35.0424 | -27.8976
17.6963 | -35.0429 | -27.9048
17.7024 | -35.0434 | -27.9123
17.7085 | -35.0439 | -27.9197
17.7146 | -35.0444 | -27.9271
17.7207 | -35.0449 | -27.9346
17.7268 | -35.0455 | -27.942
17.7329 | -35.0459 | -27.9494
17.739 | -35.0464 | -27.957
17.7451 | -35.0469 | -27.9644
17.7512 | -35.0474 | -27.9719
17.7573 | -35.0479 | -27.9795
17.7634 | -35.0484 | -27.9868
17.7695 | -35.0489 | -27.9942
17.7756 | -35.0494 | -28.0016
17.7817 | -35.0498 | -28.009
17.7878 | -35.0503 | -28.0165
17.7939 | -35.0508 | -28.0241
17.8 | -35.0513 | -28.0314
17.8061 | -35.0518 | -28.0389
17.8122 | -35.0523 | -28.0464
17.8183 | -35.0528 | -28.0538
17.8244 | -35.0533 | -28.0612
17.8305 | -35.0538 | -28.0685
17.8366 | -35.0543 | -28.0757
17.8428 | -35.0548 | -28.0832
17.8489 | -35.0553 | -28.0906
17.855 | -35.0558 | -28.0979
17.8611 | -35.0562 | -28.1052
17.8672 | -35.0568 | -28.1124
17.8733 | -35.0573 | -28.12
17.8794 | -35.0578 | -28.1274
17.8855 | -35.0582 | -28.1348
17.8916 | -35.0587 | -28.1423
17.8977 | -35.0592 | -28.1498
17.9038 | -35.0597 | -28.1569
17.9099 | -35.0602 | -28.1644
17.916 | -35.0607 | -28.1719
17.9221 | -35.0612 | -28.1793
17.9282 | -35.0617 | -28.1868
17.9343 | -35.0622 | -28.1944
17.9404 | -35.0627 | -28.2016
17.9465 | -35.0632 | -28.2091
17.9526 | -35.0637 | -28.2165
17.9587 | -35.0642 | -28.224
17.9648 | -35.0647 | -28.2313
17.9709 | -35.0652 | -28.2387
17.977 | -35.0657 | -28.246
17.9832 | -35.0661 | -28.2535
17.9893 | -35.0666 | -28.261
17.9954 | -35.0671 | -28.2683
18.0015 | -35.0676 | -28.2758
18.0076 | -35.0681 | -28.283
18.0137 | -35.0686 | -28.2904
18.0198 | -35.0691 | -28.2979
18.0259 | -35.0696 | -28.3055
18.032 | -35.07 | -28.3128
18.0381 | -35.0705 | -28.3201
18.0442 | -35.071 | -28.3274
18.0503 | -35.0715 | -28.3349
18.0564 | -35.072 | -28.3422
18.0625 | -35.0725 | -28.3498
18.0686 | -35.073 | -28.3571
18.0747 | -35.0735 | -28.3647
18.0808 | -35.074 | -28.3719
18.0869 | -35.0745 | -28.3794
18.093 | -35.075 | -28.3869
18.0991 | -35.0754 | -28.3944
18.1052 | -35.0759 | -28.4018
18.1113 | -35.0764 | -28.4092
18.1174 | -35.0769 | -28.4164
18.1236 | -35.0774 | -28.4238
18.1297 | -35.0779 | -28.4311
18.1358 | -35.0784 | -28.4385
18.1419 | -35.0789 | -28.4459
18.148 | -35.0794 | -28.4533
18.1541 | -35.0799 | -28.4604
18.1602 | -35.0804 | -28.4678
18.1663 | -35.0809 | -28.4752
18.1724 | -35.0813 | -28.4825
18.1785 | -35.0818 | -28.4899
18.1846 | -35.0823 | -28.4973
18.1907 | -35.0828 | -28.5047
18.1968 | -35.0833 | -28.5119
18.2029 | -35.0838 | -28.5192
18.209 | -35.0843 | -28.5268
18.2151 | -35.0848 | -28.5343
18.2212 | -35.0853 | -28.5413
18.2273 | -35.0858 | -28.5487
18.2334 | -35.0862 | -28.5561
18.2395 | -35.0867 | -28.5634
18.2456 | -35.0872 | -28.5707
18.2517 | -35.0877 | -28.5782
18.2578 | -35.0882 | -28.5853
18.2639 | -35.0887 | -28.5927
18.2701 | -35.0892 | -28.5999
18.2762 | -35.0897 | -28.6072
18.2823 | -35.0902 | -28.6144
18.2884 | -35.0907 | -28.6217
18.2945 | -35.0912 | -28.6286
18.3006 | -35.0917 | -28.6361
18.3067 | -35.0922 | -28.6434
18.3128 | -35.0927 | -28.6508
18.3189 | -35.0932 | -28.6582
18.325 | -35.0937 | -28.6654
18.3311 | -35.0942 | -28.6729
18.3372 | -35.0946 | -28.6803
18.3433 | -35.0951 | -28.6876
18.3494 | -35.0956 | -28.6951
18.3555 | -35.0961 | -28.7025
18.3616 | -35.0966 | -28.7094
18.3677 | -35.0971 | -28.7166
18.3738 | -35.0976 | -28.7239
18.3799 | -35.0981 | -28.7313
18.386 | -35.0986 | -28.7387
18.3921 | -35.0991 | -28.746
18.3982 | -35.0996 | -28.753
18.4043 | -35.1001 | -28.7603
18.4105 | -35.1006 | -28.7677
18.4166 | -35.101 | -28.7749
18.4227 | -35.1015 | -28.7822
18.4288 | -35.102 | -28.7896
18.4349 | -35.1024 | -28.7967
18.441 | -35.1029 | -28.8039
18.4471 | -35.1034 | -28.8112
18.4532 | -35.1039 | -28.8187
18.4593 | -35.1044 | -28.8258
18.4654 | -35.1049 | -28.8328
18.4715 | -35.1054 | -28.8402
18.4776 | -35.1059 | -28.8477
18.4837 | -35.1063 | -28.855
18.4898 | -35.1068 | -28.8623
18.4959 | -35.1073 | -28.8698
18.502 | -35.1078 | -28.877
18.5081 | -35.1083 | -28.8843
18.5142 | -35.1087 | -28.8917
18.5203 | -35.1092 | -28.8989
18.5264 | -35.1097 | -28.906
18.5325 | -35.1102 | -28.9134
18.5386 | -35.1107 | -28.9205
18.5447 | -35.1112 | -28.9278
18.5508 | -35.1117 | -28.9351
18.557 | -35.1122 | -28.9425
18.5631 | -35.1127 | -28.9498
18.5692 | -35.1131 | -28.9571
18.5753 | -35.1136 | -28.9641
18.5814 | -35.1141 | -28.9714
18.5875 | -35.1146 | -28.9785
18.5936 | -35.1151 | -28.9857
18.5997 | -35.1155 | -28.9929
18.6058 | -35.116 | -29.0001
18.6119 | -35.1165 | -29.0073
18.618 | -35.117 | -29.0146
18.6241 | -35.1175 | -29.0218
18.6302 | -35.118 | -29.0292
18.6363 | -35.1185 | -29.0364
18.6424 | -35.119 | -29.0434
18.6485 | -35.1195 | -29.0506
18.6546 | -35.1199 | -29.0579
18.6607 | -35.1204 | -29.0651
18.6668 | -35.1209 | -29.0725
18.6729 | -35.1214 | -29.0798
18.679 | -35.1219 | -29.0869
18.6851 | -35.1224 | -29.094
18.6912 | -35.1229 | -29.1012
18.6974 | -35.1234 | -29.1084
18.7035 | -35.1239 | -29.1158
18.7096 | -35.1244 | -29.1231
18.7157 | -35.1248 | -29.1302
18.7218 | -35.1253 | -29.1374
18.7279 | -35.1258 | -29.1448
18.734 | -35.1263 | -29.1522
18.7401 | -35.1268 | -29.1594
18.7462 | -35.1273 | -29.1663
18.7523 | -35.1278 | -29.1736
18.7584 | -35.1283 | -29.1809
18.7645 | -35.1288 | -29.188
18.7706 | -35.1292 | -29.1951
18.7767 | -35.1297 | -29.2023
18.7828 | -35.1302 | -29.2093
18.7889 | -35.1307 | -29.2166
18.795 | -35.1312 | -29.2238
18.8011 | -35.1316 | -29.231
18.8072 | -35.1321 | -29.2384
18.8133 | -35.1326 | -29.2455
18.8194 | -35.1331 | -29.2526
18.8255 | -35.1336 | -29.2598
18.8316 | -35.134 | -29.2671
18.8377 | -35.1345 | -29.2742
18.8439 | -35.135 | -29.2814
18.85 | -35.1355 | -29.2887
18.8561 | -35.136 | -29.2957
18.8622 | -35.1365 | -29.3029
18.8683 | -35.1369 | -29.31
18.8744 | -35.1374 | -29.3173
18.8805 | -35.1379 | -29.3247
18.8866 | -35.1384 | -29.332
18.8927 | -35.1389 | -29.339
18.8988 | -35.1394 | -29.3462
18.9049 | -35.1399 | -29.3533
18.911 | -35.1404 | -29.3604
18.9171 | -35.1409 | -29.3675
18.9232 | -35.1414 | -29.3744
18.9293 | -35.1419 | -29.3815
18.9354 | -35.1424 | -29.3887
18.9415 | -35.1429 | -29.3961
18.9476 | -35.1433 | -29.4032
18.9537 | -35.1438 | -29.4104
18.9598 | -35.1443 | -29.4172
18.9659 | -35.1448 | -29.4245
18.972 | -35.1452 | -29.4317
18.9781 | -35.1457 | -29.4389
18.9843 | -35.1462 | -29.4462
18.9904 | -35.1467 | -29.4532
18.9965 | -35.1471 | -29.4602
19.0026 | -35.1476 | -29.4674
19.0087 | -35.1481 | -29.4748
19.0148 | -35.1486 | -29.4819
19.0209 | -35.1491 | -29.489
19.027 | -35.1496 | -29.4962
19.0331 | -35.15 | -29.5031
19.0392 | -35.1505 | -29.5104
19.0453 | -35.151 | -29.5176
19.0514 | -35.1515 | -29.5247
19.0575 | -35.152 | -29.532
19.0636 | -35.1525 | -29.5388
19.0697 | -35.1529 | -29.5459
19.0758 | -35.1534 | -29.5531
19.0819 | -35.1539 | -29.5602
19.088 | -35.1544 | -29.5673
19.0941 | -35.1549 | -29.5746
19.1002 | -35.1554 | -29.5815
19.1063 | -35.1558 | -29.5886
19.1124 | -35.1563 | -29.5958
19.1185 | -35.1568 | -29.6029
19.1246 | -35.1573 | -29.6099
19.1308 | -35.1577 | -29.6171
19.1369 | -35.1582 | -29.6239
19.143 | -35.1587 | -29.6309
19.1491 | -35.1592 | -29.6381
19.1552 | -35.1596 | -29.6453
19.1613 | -35.1601 | -29.6523
19.1674 | -35.1606 | -29.6595
19.1735 | -35.161 | -29.6663
19.1796 | -35.1615 | -29.6732
19.1857 | -35.162 | -29.6803
19.1918 | -35.1625 | -29.6877
19.1979 | -35.163 | -29.695
19.204 | -35.1635 | -29.7016
19.2101 | -35.1639 | -29.7086
19.2162 | -35.1644 | -29.7157
19.2223 | -35.1649 | -29.7228
19.2284 | -35.1653 | -29.73
19.2345 | -35.1658 | -29.737
19.2406 | -35.1662 | -29.7439
19.2467 | -35.1667 | -29.751
19.2528 | -35.1672 | -29.7582
19.2589 | -35.1676 | -29.765
19.265 | -35.1681 | -29.7723
19.2712 | -35.1686 | -29.7793
19.2773 | -35.1691 | -29.7863
19.2834 | -35.1695 | -29.7932
19.2895 | -35.17 | -29.8004
19.2956 | -35.1705 | -29.8077
19.3017 | -35.1709 | -29.8149
19.3078 | -35.1714 | -29.822
19.3139 | -35.1719 | -29.829
19.32 | -35.1723 | -29.836
19.3261 | -35.1728 | -29.8432
19.3322 | -35.1733 | -29.8502
19.3383 | -35.1738 | -29.8573
19.3444 | -35.1743 | -29.8641
19.3505 | -35.1747 | -29.8711
19.3566 | -35.1752 | -29.8782
19.3627 | -35.1756 | -29.8852
19.3688 | -35.1761 | -29.8922
19.3749 | -35.1766 | -29.8993
19.381 | -35.1771 | -29.9061
19.3871 | -35.1776 | -29.9131
19.3932 | -35.178 | -29.9201
19.3993 | -35.1785 | -29.9271
19.4054 | -35.179 | -29.9342
19.4115 | -35.1794 | -29.9415
19.4177 | -35.1799 | -29.9485
19.4238 | -35.1804 | -29.9555
19.4299 | -35.1809 | -29.9625
19.436 | -35.1813 | -29.9697
19.4421 | -35.1818 | -29.9767
19.4482 | -35.1823 | -29.9838
19.4543 | -35.1827 | -29.9904
19.4604 | -35.1832 | -29.9975
19.4665 | -35.1837 | -30.0046
19.4726 | -35.1841 | -30.0115
19.4787 | -35.1846 | -30.0184
19.4848 | -35.185 | -30.0254
19.4909 | -35.1855 | -30.0323
19.497 | -35.1859 | -30.0392
19.5031 | -35.1864 | -30.0462
19.5092 | -35.1868 | -30.0532
19.5153 | -35.1873 | -30.0602
19.5214 | -35.1878 | -30.0669
19.5275 | -35.1882 | -30.074
19.5336 | -35.1887 | -30.081
19.5397 | -35.1891 | -30.0879
19.5458 | -35.1896 | -30.095
19.5519 | -35.1901 | -30.102
19.5581 | -35.1906 | -30.1087
19.5642 | -35.191 | -30.1158
19.5703 | -35.1915 | -30.1227
19.5764 | -35.1919 | -30.1299
19.5825 | -35.1924 | -30.1366
19.5886 | -35.1929 | -30.1435
19.5947 | -35.1934 | -30.1504
19.6008 | -35.1938 | -30.1573
19.6069 | -35.1943 | -30.1642
19.613 | -35.1947 | -30.1711
19.6191 | -35.1952 | -30.1781
19.6252 | -35.1957 | -30.1852
19.6313 | -35.1962 | -30.1919
19.6374 | -35.1967 | -30.1988
19.6435 | -35.1971 | -30.2056
19.6496 | -35.1976 | -30.2126
19.6557 | -35.1981 | -30.2196
19.6618 | -35.1986 | -30.2264
19.6679 | -35.1991 | -30.2333
19.674 | -35.1995 | -30.2402
19.6801 | -35.1999 | -30.2471
19.6862 | -35.2004 | -30.254
19.6923 | -35.2009 | -30.261
19.6984 | -35.2013 | -30.268
19.7046 | -35.2018 | -30.2749
19.7107 | -35.2023 | -30.2819
19.7168 | -35.2027 | -30.2887
19.7229 | -35.2032 | -30.2955
19.729 | -35.2036 | -30.3023
19.7351 | -35.2041 | -30.3088
19.7412 | -35.2046 | -30.3158
19.7473 | -35.2051 | -30.3226
19.7534 | -35.2055 | -30.3295
19.7595 | -35.206 | -30.3363
19.7656 | -35.2064 | -30.3432
19.7717 | -35.2069 | -30.3499
19.7778 | -35.2074 | -30.3566
19.7839 | -35.2079 | -30.3636
19.79 | -35.2083 | -30.3704
19.7961 | -35.2088 | -30.3775
19.8022 | -35.2092 | -30.3844
19.8083 | -35.2097 | -30.3913
19.8144 | -35.2102 | -30.3981
19.8205 | -35.2106 | -30.4051
19.8266 | -35.211 | -30.412
19.8327 | -35.2115 | -30.419
19.8388 | -35.212 | -30.4254
19.845 | -35.2124 | -30.4324
19.8511 | -35.2128 | -30.4395
19.8572 | -35.2133 | -30.4462
19.8633 | -35.2137 | -30.4531
19.8694 | -35.2142 | -30.4602
19.8755 | -35.2146 | -30.467
19.8816 | -35.2151 | -30.4739
19.8877 | -35.2156 | -30.4809
19.8938 | -35.2161 | -30.4877
19.8999 | -35.2165 | -30.4947
19.906 | -35.217 | -30.5016
19.9121 | -35.2174 | -30.5081
19.9182 | -35.2179 | -30.5148
19.9243 | -35.2184 | -30.5215
19.9304 | -35.2188 | -30.5282
19.9365 | -35.2193 | -30.5353
19.9426 | -35.2197 | -30.542
19.9487 | -35.2202 | -30.549
19.9548 | -35.2207 | -30.5558
19.9609 | -35.2211 | -30.5626
19.967 | -35.2216 | -30.5694
19.9731 | -35.2221 | -30.5763
19.9792 | -35.2225 | -30.5829
19.9853 | -35.223 | -30.5897
19.9915 | -35.2235 | -30.5965
19.9976 | -35.2239 | -30.6033
20.0037 | -35.2244 | -30.6103
20.0098 | -35.2248 | -30.6171
20.0159 | -35.2253 | -30.6239
20.022 | -35.2257 | -30.6308
20.0281 | -35.2262 | -30.6375
20.0342 | -35.2266 | -30.6445
20.0403 | -35.2271 | -30.6515
20.0464 | -35.2275 | -30.6584
20.0525 | -35.228 | -30.6653
20.0586 | -35.2284 | -30.6722
20.0647 | -35.2289 | -30.6791
20.0708 | -35.2294 | -30.6862
20.0769 | -35.2298 | -30.693
20.083 | -35.2303 | -30.6997
20.0891 | -35.2308 | -30.7064
20.0952 | -35.2313 | -30.7131
20.1013 | -35.2317 | -30.7202
20.1074 | -35.2322 | -30.727
20.1135 | -35.2326 | -30.734
20.1196 | -35.2331 | -30.7406
20.1257 | -35.2335 | -30.7473
20.1319 | -35.234 | -30.7542
20.138 | -35.2344 | -30.761
20.1441 | -35.2349 | -30.7679
20.1502 | -35.2354 | -30.7748
20.1563 | -35.2358 | -30.7813
20.1624 | -35.2363 | -30.7879
20.1685 | -35.2367 | -30.7946
20.1746 | -35.2372 | -30.8014
20.1807 | -35.2376 | -30.8082
20.1868 | -35.2381 | -30.8151
20.1929 | -35.2385 | -30.8219
20.199 | -35.239 | -30.8287
20.2051 | -35.2394 | -30.8354
20.2112 | -35.2399 | -30.8421
20.2173 | -35.2403 | -30.8489
20.2234 | -35.2408 | -30.8556
20.2295 | -35.2412 | -30.8623
20.2356 | -35.2417 | -30.869
20.2417 | -35.2421 | -30.8759
20.2478 | -35.2425 | -30.8827
20.2539 | -35.243 | -30.8898
20.26 | -35.2434 | -30.8963
20.2661 | -35.2439 | -30.9032
20.2723 | -35.2444 | -30.9098
20.2784 | -35.2448 | -30.9164
20.2845 | -35.2452 | -30.923
20.2906 | -35.2457 | -30.9296
20.2967 | -35.2461 | -30.9361
20.3028 | -35.2466 | -30.9429
20.3089 | -35.2471 | -30.9497
20.315 | -35.2476 | -30.9563
20.3211 | -35.248 | -30.9629
20.3272 | -35.2485 | -30.9696
20.3333 | -35.249 | -30.9761
20.3394 | -35.2494 | -30.9826
20.3455 | -35.2499 | -30.9894
20.3516 | -35.2503 | -30.9962
20.3577 | -35.2507 | -31.0031
20.3638 | -35.2512 | -31.0098
20.3699 | -35.2517 | -31.0166
20.376 | -35.2521 | -31.0233
20.3821 | -35.2526 | -31.0301
20.3882 | -35.253 | -31.0368
20.3943 | -35.2534 | -31.0436
20.4004 | -35.2539 | -31.0499
20.4065 | -35.2543 | -31.0566
20.4126 | -35.2548 | -31.0636
20.4188 | -35.2552 | -31.0706
20.4249 | -35.2557 | -31.0773
20.431 | -35.2561 | -31.0841
20.4371 | -35.2566 | -31.0907
20.4432 | -35.257 | -31.0974
20.4493 | -35.2575 | -31.1043
20.4554 | -35.258 | -31.111
20.4615 | -35.2585 | -31.1177
20.4676 | -35.2589 | -31.1245
20.4737 | -35.2593 | -31.1311
20.4798 | -35.2598 | -31.1379
20.4859 | -35.2603 | -31.1447
20.492 | -35.2607 | -31.1512
20.4981 | -35.2612 | -31.1579
20.5042 | -35.2616 | -31.1647
20.5103 | -35.2621 | -31.1714
20.5164 | -35.2625 | -31.178
20.5225 | -35.263 | -31.1848
20.5286 | -35.2635 | -31.1916
20.5347 | -35.2639 | -31.1983
20.5408 | -35.2644 | -31.2049
20.5469 | -35.2648 | -31.2117
20.553 | -35.2653 | -31.2185
20.5592 | -35.2657 | -31.2252
20.5653 | -35.2662 | -31.2317
20.5714 | -35.2666 | -31.2383
20.5775 | -35.2671 | -31.2449
20.5836 | -35.2676 | -31.2517
20.5897 | -35.268 | -31.2586
20.5958 | -35.2685 | -31.2655
20.6019 | -35.2689 | -31.2721
20.608 | -35.2694 | -31.2788
20.6141 | -35.2698 | -31.2852
20.6202 | -35.2703 | -31.2918
20.6263 | -35.2708 | -31.2986
20.6324 | -35.2712 | -31.3051
20.6385 | -35.2717 | -31.3118
20.6446 | -35.2721 | -31.3185
20.6507 | -35.2726 | -31.325
20.6568 | -35.273 | -31.3317
20.6629 | -35.2735 | -31.3386
20.669 | -35.2739 | -31.3454
20.6751 | -35.2744 | -31.3519
20.6812 | -35.2749 | -31.3585
20.6873 | -35.2753 | -31.3653
20.6934 | -35.2757 | -31.372
20.6995 | -35.2762 | -31.3785
20.7057 | -35.2767 | -31.3851
20.7118 | -35.2771 | -31.3917
20.7179 | -35.2776 | -31.3981
20.724 | -35.278 | -31.405
20.7301 | -35.2785 | -31.4119
20.7362 | -35.2789 | -31.4183
20.7423 | -35.2793 | -31.425
20.7484 | -35.2798 | -31.4315
20.7545 | -35.2802 | -31.438
20.7606 | -35.2807 | -31.4447
20.7667 | -35.2811 | -31.4514
20.7728 | -35.2816 | -31.458
20.7789 | -35.282 | -31.4645
20.785 | -35.2825 | -31.4711
20.7911 | -35.2829 | -31.4772
20.7972 | -35.2834 | -31.4839
20.8033 | -35.2838 | -31.4906
20.8094 | -35.2843 | -31.4973
20.8155 | -35.2848 | -31.5036
20.8216 | -35.2852 | -31.5103
20.8277 | -35.2856 | -31.5166
20.8338 | -35.2861 | -31.5231
20.8399 | -35.2866 | -31.5298
20.8461 | -35.287 | -31.5365
20.8522 | -35.2875 | -31.5432
20.8583 | -35.2879 | -31.55
20.8644 | -35.2884 | -31.5567
20.8705 | -35.2889 | -31.5635
20.8766 | -35.2893 | -31.5701
20.8827 | -35.2898 | -31.5765
20.8888 | -35.2902 | -31.5831
20.8949 | -35.2907 | -31.5893
20.901 | -35.2911 | -31.596
20.9071 | -35.2915 | -31.6024
20.9132 | -35.292 | -31.609
20.9193 | -35.2924 | -31.6155
20.9254 | -35.2928 | -31.6223
20.9315 | -35.2933 | -31.6286
20.9376 | -35.2938 | -31.6352
20.9437 | -35.2942 | -31.6418
20.9498 | -35.2947 | -31.6482
20.9559 | -35.2951 | -31.6546
20.962 | -35.2956 | -31.6611
20.9681 | -35.296 | -31.6675
20.9742 | -35.2965 | -31.6739
20.9803 | -35.2969 | -31.6805
20.9864 | -35.2974 | -31.6872
20.9926 | -35.2978 | -31.6937
20.9987 | -35.2982 | -31.7001
21.0048 | -35.2987 | -31.7066
21.0109 | -35.2992 | -31.7131
21.017 | -35.2996 | -31.7194
21.0231 | -35.3001 | -31.7259
21.0292 | -35.3005 | -31.7325
21.0353 | -35.3009 | -31.739
21.0414 | -35.3014 | -31.7454
21.0475 | -35.3019 | -31.7519
21.0536 | -35.3023 | -31.7584
21.0597 | -35.3027 | -31.765
21.0658 | -35.3032 | -31.7714
21.0719 | -35.3036 | -31.7777
21.078 | -35.3041 | -31.7845
21.0841 | -35.3045 | -31.791
21.0902 | -35.3051 | -31.7975
21.0963 | -35.3055 | -31.8042
21.1024 | -35.306 | -31.8108
21.1085 | -35.3064 | -31.8175
21.1146 | -35.3069 | -31.824
21.1207 | -35.3073 | -31.8304
21.1268 | -35.3078 | -31.837
21.133 | -35.3082 | -31.8435
21.1391 | -35.3087 | -31.8497
21.1452 | -35.3091 | -31.8561
21.1513 | -35.3096 | -31.8626
21.1574 | -35.31 | -31.8691
21.1635 | -35.3105 | -31.8757
21.1696 | -35.3109 | -31.8822
21.1757 | -35.3114 | -31.8886
21.1818 | -35.3118 | -31.8952
21.1879 | -35.3123 | -31.9019
21.194 | -35.3127 | -31.9085
21.2001 | -35.3131 | -31.915
21.2062 | -35.3136 | -31.9216
21.2123 | -35.3141 | -31.928
21.2184 | -35.3145 | -31.9345
21.2245 | -35.315 | -31.9409
21.2306 | -35.3154 | -31.9476
21.2367 | -35.3158 | -31.954
21.2428 | -35.3163 | -31.9604
21.2489 | -35.3168 | -31.967
21.255 | -35.3172 | -31.9735
21.2611 | -35.3176 | -31.9799
21.2672 | -35.3181 | -31.9864
21.2733 | -35.3185 | -31.993
21.2795 | -35.3189 | -31.9992
21.2856 | -35.3194 | -32.0056
21.2917 | -35.3198 | -32.012
21.2978 | -35.3203 | -32.0185
21.3039 | -35.3207 | -32.025
21.31 | -35.3212 | -32.0316
21.3161 | -35.3216 | -32.0379
21.3222 | -35.3221 | -32.0443
21.3283 | -35.3225 | -32.0506
21.3344 | -35.323 | -32.0568
21.3405 | -35.3234 | -32.0634
21.3466 | -35.3239 | -32.0698
21.3527 | -35.3243 | -32.0762
21.3588 | -35.3248 | -32.0826
21.3649 | -35.3252 | -32.089
21.371 | -35.3257 | -32.0956
21.3771 | -35.3261 | -32.102
21.3832 | -35.3266 | -32.1086
21.3893 | -35.327 | -32.1148
21.3954 | -35.3275 | -32.1213
21.4015 | -35.3279 | -32.1279
21.4076 | -35.3284 | -32.1344
21.4137 | -35.3289 | -32.1409
21.4199 | -35.3293 | -32.1476
21.426 | -35.3297 | -32.1537
21.4321 | -35.3302 | -32.1602
21.4382 | -35.3306 | -32.1667
21.4443 | -35.331 | -32.1731
21.4504 | -35.3315 | -32.1797
21.4565 | -35.332 | -32.1858
21.4626 | -35.3324 | -32.1924
21.4687 | -35.3328 | -32.199
21.4748 | -35.3333 | -32.2054
21.4809 | -35.3337 | -32.2118
21.487 | -35.3342 | -32.2183
21.4931 | -35.3346 | -32.2245
21.4992 | -35.3351 | -32.2308
21.5053 | -35.3355 | -32.2374
21.5114 | -35.336 | -32.2438
21.5175 | -35.3364 | -32.2502
21.5236 | -35.3369 | -32.2565
21.5297 | -35.3373 | -32.2628
21.5358 | -35.3377 | -32.2691
21.5419 | -35.3382 | -32.2755
21.548 | -35.3386 | -32.282
21.5541 | -35.3391 | -32.2887
21.5602 | -35.3395 | -32.2953
21.5664 | -35.34 | -32.3016
21.5725 | -35.3404 | -32.308
21.5786 | -35.3408 | -32.3145
21.5847 | -35.3413 | -32.3212
21.5908 | -35.3418 | -32.3278
21.5969 | -35.3422 | -32.334
21.603 | -35.3427 | -32.3406
21.6091 | -35.3431 | -32.347
21.6152 | -35.3435 | -32.3537
21.6213 | -35.3439 | -32.3602
21.6274 | -35.3444 | -32.3665
21.6335 | -35.3448 | -32.3728
21.6396 | -35.3453 | -32.3793
21.6457 | -35.3457 | -32.3859
21.6518 | -35.3462 | -32.3923
21.6579 | -35.3466 | -32.3986
21.664 | -35.3471 | -32.4051
21.6701 | -35.3475 | -32.4114
21.6762 | -35.348 | -32.4178
21.6823 | -35.3484 | -32.4244
21.6884 | -35.3488 | -32.4311
21.6945 | -35.3493 | -32.4374
21.7006 | -35.3498 | -32.444
21.7068 | -35.3502 | -32.4503
21.7129 | -35.3507 | -32.4567
21.719 | -35.3511 | -32.463
21.7251 | -35.3516 | -32.4694
21.7312 | -35.352 | -32.4757
21.7373 | -35.3525 | -32.4821
21.7434 | -35.353 | -32.4885
21.7495 | -35.3534 | -32.4949
21.7556 | -35.3539 | -32.5013
21.7617 | -35.3543 | -32.5077
21.7678 | -35.3548 | -32.5141
21.7739 | -35.3553 | -32.5204
21.78 | -35.3557 | -32.5266
21.7861 | -35.3562 | -32.5332
21.7922 | -35.3566 | -32.5396
21.7983 | -35.357 | -32.5458
21.8044 | -35.3575 | -32.5522
21.8105 | -35.358 | -32.5584
21.8166 | -35.3584 | -32.5648
21.8227 | -35.3589 | -32.5711
21.8288 | -35.3594 | -32.5776
21.8349 | -35.3598 | -32.584
21.841 | -35.3602 | -32.5904
21.8471 | -35.3607 | -32.5966
21.8533 | -35.3612 | -32.603
21.8594 | -35.3616 | -32.6094
21.8655 | -35.362 | -32.6156
21.8716 | -35.3625 | -32.6222
21.8777 | -35.3629 | -32.6285
21.8838 | -35.3633 | -32.6349
21.8899 | -35.3638 | -32.6413
21.896 | -35.3642 | -32.6477
21.9021 | -35.3647 | -32.6541
21.9082 | -35.3651 | -32.6606
21.9143 | -35.3655 | -32.6668
21.9204 | -35.366 | -32.673
21.9265 | -35.3664 | -32.6793
21.9326 | -35.3669 | -32.6857
21.9387 | -35.3674 | -32.6919
21.9448 | -35.3678 | -32.6983
21.9509 | -35.3683 | -32.7044
21.957 | -35.3687 | -32.7108
21.9631 | -35.3692 | -32.7173
21.9692 | -35.3696 | -32.7236
21.9753 | -35.3701 | -32.7299
21.9814 | -35.3705 | -32.7363
21.9875 | -35.371 | -32.7423
21.9937 | -35.3714 | -32.7488
21.9998 | -35.3719 | -32.7551
22.0059 | -35.3724 | -32.7615
22.012 | -35.3728 | -32.7678
22.0181 | -35.3733 | -32.774
22.0242 | -35.3738 | -32.7803
22.0303 | -35.3742 | -32.7868
22.0364 | -35.3746 | -32.7929
22.0425 | -35.3751 | -32.7992
22.0486 | -35.3755 | -32.8055
22.0547 | -35.3759 | -32.8118
22.0608 | -35.3764 | -32.8182
22.0669 | -35.3768 | -32.8246
22.073 | -35.3772 | -32.8308
22.0791 | -35.3777 | -32.8372
22.0852 | -35.3781 | -32.8436
22.0913 | -35.3786 | -32.85
22.0974 | -35.3791 | -32.8564
22.1035 | -35.3795 | -32.8628
22.1096 | -35.38 | -32.8694
22.1157 | -35.3804 | -32.8758
22.1218 | -35.3809 | -32.8822
22.1279 | -35.3813 | -32.8883
22.134 | -35.3818 | -32.8946
22.1402 | -35.3822 | -32.9009
22.1463 | -35.3827 | -32.9072
22.1524 | -35.3832 | -32.9135
22.1585 | -35.3836 | -32.9199
22.1646 | -35.3841 | -32.9258
22.1707 | -35.3845 | -32.9323
22.1768 | -35.385 | -32.9386
22.1829 | -35.3854 | -32.9449
22.189 | -35.3859 | -32.9512
22.1951 | -35.3863 | -32.9575
22.2012 | -35.3868 | -32.9637
22.2073 | -35.3872 | -32.9701
22.2134 | -35.3877 | -32.9763
22.2195 | -35.3881 | -32.9826
22.2256 | -35.3886 | -32.9888
22.2317 | -35.389 | -32.9949
22.2378 | -35.3895 | -33.0013
22.2439 | -35.3899 | -33.0077
22.25 | -35.3904 | -33.0139
22.2561 | -35.3908 | -33.0203
22.2622 | -35.3912 | -33.0267
22.2683 | -35.3917 | -33.0328
22.2744 | -35.3921 | -33.0391
22.2806 | -35.3926 | -33.0456
22.2867 | -35.393 | -33.0517
22.2928 | -35.3935 | -33.0579
22.2989 | -35.3939 | -33.0642
22.305 | -35.3944 | -33.0704
22.3111 | -35.3949 | -33.0768
22.3172 | -35.3953 | -33.0832
22.3233 | -35.3958 | -33.0897
22.3294 | -35.3962 | -33.096
22.3355 | -35.3966 | -33.1023
22.3416 | -35.3971 | -33.1086
22.3477 | -35.3975 | -33.115
22.3538 | -35.398 | -33.1212
22.3599 | -35.3984 | -33.1273
22.366 | -35.3988 | -33.1335
22.3721 | -35.3993 | -33.1397
22.3782 | -35.3998 | -33.146
22.3843 | -35.4002 | -33.1523
22.3904 | -35.4007 | -33.1584
22.3965 | -35.4011 | -33.1647
22.4026 | -35.4015 | -33.171
22.4087 | -35.402 | -33.1768
22.4148 | -35.4025 | -33.1831
22.4209 | -35.4029 | -33.1895
22.4271 | -35.4033 | -33.1957
22.4332 | -35.4038 | -33.202
22.4393 | -35.4042 | -33.2084
22.4454 | -35.4047 | -33.2143
22.4515 | -35.4051 | -33.2206
22.4576 | -35.4056 | -33.2268
22.4637 | -35.406 | -33.2332
22.4698 | -35.4065 | -33.2395
22.4759 | -35.407 | -33.2457
22.482 | -35.4074 | -33.2521
22.4881 | -35.4079 | -33.2584
22.4942 | -35.4083 | -33.2647
22.5003 | -35.4088 | -33.271
22.5064 | -35.4092 | -33.2774
22.5125 | -35.4097 | -33.2835
22.5186 | -35.4101 | -33.2897
22.5247 | -35.4106 | -33.296
22.5308 | -35.411 | -33.3023
22.5369 | -35.4115 | -33.3084
22.543 | -35.4119 | -33.3145
22.5491 | -35.4124 | -33.3206
22.5552 | -35.4128 | -33.3269
22.5613 | -35.4133 | -33.3331
22.5675 | -35.4138 | -33.3395
22.5736 | -35.4142 | -33.3458
22.5797 | -35.4147 | -33.3521
22.5858 | -35.4152 | -33.3582
22.5919 | -35.4156 | -33.3644
22.598 | -35.4161 | -33.3707
22.6041 | -35.4165 | -33.3769
22.6102 | -35.417 | -33.3832
22.6163 | -35.4174 | -33.3893
22.6224 | -35.4179 | -33.3956
22.6285 | -35.4183 | -33.4017
22.6346 | -35.4188 | -33.408
22.6407 | -35.4192 | -33.4142
22.6468 | -35.4197 | -33.4203
22.6529 | -35.4201 | -33.4263
22.659 | -35.4206 | -33.4325
22.6651 | -35.421 | -33.4387
22.6712 | -35.4215 | -33.4449
22.6773 | -35.4219 | -33.4512
22.6834 | -35.4224 | -33.4577
22.6895 | -35.4228 | -33.4636
22.6956 | -35.4233 | -33.4697
22.7017 | -35.4237 | -33.4759
22.7079 | -35.4242 | -33.4819
22.714 | -35.4246 | -33.4882
22.7201 | -35.4251 | -33.4943
22.7262 | -35.4255 | -33.5005
22.7323 | -35.426 | -33.5067
22.7384 | -35.4264 | -33.5129
22.7445 | -35.4269 | -33.5192
22.7506 | -35.4273 | -33.5254
22.7567 | -35.4277 | -33.5315
22.7628 | -35.4282 | -33.5375
22.7689 | -35.4286 | -33.5437
22.775 | -35.4291 | -33.5499
22.7811 | -35.4296 | -33.5562
22.7872 | -35.43 | -33.5624
22.7933 | -35.4305 | -33.5683
22.7994 | -35.431 | -33.5746
22.8055 | -35.4314 | -33.5808
22.8116 | -35.4319 | -33.5869
22.8177 | -35.4323 | -33.593
22.8238 | -35.4328 | -33.5994
22.8299 | -35.4333 | -33.6054
22.836 | -35.4337 | -33.6117
22.8421 | -35.4342 | -33.618
22.8482 | -35.4346 | -33.6242
22.8544 | -35.4351 | -33.6304
22.8605 | -35.4355 | -33.6365
22.8666 | -35.436 | -33.6426
22.8727 | -35.4364 | -33.6487
22.8788 | -35.4368 | -33.6549
22.8849 | -35.4373 | -33.6611
22.891 | -35.4377 | -33.6673
22.8971 | -35.4381 | -33.6736
22.9032 | -35.4386 | -33.6796
22.9093 | -35.4391 | -33.6858
22.9154 | -35.4395 | -33.692
22.9215 | -35.4399 | -33.6981
22.9276 | -35.4404 | -33.7043
22.9337 | -35.4408 | -33.7103
22.9398 | -35.4413 | -33.7164
22.9459 | -35.4417 | -33.7226
22.952 | -35.4422 | -33.7288
22.9581 | -35.4426 | -33.735
22.9642 | -35.4431 | -33.7411
22.9703 | -35.4435 | -33.7471
22.9764 | -35.444 | -33.7534
22.9825 | -35.4444 | -33.7597
22.9886 | -35.4449 | -33.7659
22.9948 | -35.4453 | -33.7721
23.0009 | -35.4458 | -33.7784
23.007 | -35.4463 | -33.7844
23.0131 | -35.4467 | -33.7906
23.0192 | -35.4471 | -33.7968
23.0253 | -35.4476 | -33.8029
23.0314 | -35.448 | -33.8091
23.0375 | -35.4485 | -33.8152
23.0436 | -35.4489 | -33.8211
23.0497 | -35.4494 | -33.8273
23.0558 | -35.4498 | -33.8336
23.0619 | -35.4502 | -33.8398
23.068 | -35.4507 | -33.8461
23.0741 | -35.4511 | -33.852
23.0802 | -35.4516 | -33.8582
23.0863 | -35.452 | -33.8644
23.0924 | -35.4524 | -33.8705
23.0985 | -35.4529 | -33.8767
23.1046 | -35.4533 | -33.8829
23.1107 | -35.4538 | -33.8891
23.1168 | -35.4542 | -33.8954
23.1229 | -35.4546 | -33.9014
23.129 | -35.4551 | -33.9076
23.1351 | -35.4555 | -33.9136
23.1413 | -35.456 | -33.9199
23.1474 | -35.4564 | -33.9258
23.1535 | -35.4568 | -33.9319
23.1596 | -35.4573 | -33.938
23.1657 | -35.4577 | -33.9441
23.1718 | -35.4582 | -33.9502
23.1779 | -35.4586 | -33.9563
23.184 | -35.459 | -33.9622
23.1901 | -35.4595 | -33.9684
23.1962 | -35.46 | -33.9745
23.2023 | -35.4604 | -33.9806
23.2084 | -35.4609 | -33.9867
23.2145 | -35.4613 | -33.9927
23.2206 | -35.4617 | -33.9989
23.2267 | -35.4622 | -34.0051
23.2328 | -35.4626 | -34.0113
23.2389 | -35.4631 | -34.0173
23.245 | -35.4635 | -34.0235
23.2511 | -35.464 | -34.0293
23.2572 | -35.4644 | -34.0355
23.2633 | -35.4649 | -34.0417
23.2694 | -35.4653 | -34.0478
23.2755 | -35.4657 | -34.0539
23.2817 | -35.4662 | -34.0599
23.2878 | -35.4666 | -34.0659
23.2939 | -35.4671 | -34.0721
23.3 | -35.4675 | -34.0783
23.3061 | -35.468 | -34.0843
23.3122 | -35.4684 | -34.0903
23.3183 | -35.4689 | -34.0964
23.3244 | -35.4693 | -34.1024
23.3305 | -35.4698 | -34.1085
23.3366 | -35.4702 | -34.1146
23.3427 | -35.4706 | -34.1207
23.3488 | -35.4711 | -34.1268
23.3549 | -35.4715 | -34.1329
23.361 | -35.472 | -34.1389
23.3671 | -35.4724 | -34.1451
23.3732 | -35.4728 | -34.1511
23.3793 | -35.4733 | -34.1572
23.3854 | -35.4737 | -34.1633
23.3915 | -35.4742 | -34.1693
23.3976 | -35.4746 | -34.1754
23.4037 | -35.4751 | -34.1812
23.4098 | -35.4755 | -34.1874
23.4159 | -35.476 | -34.1935
23.422 | -35.4764 | -34.1996
23.4282 | -35.4769 | -34.2054
23.4343 | -35.4773 | -34.2115
23.4404 | -35.4778 | -34.2176
23.4465 | -35.4782 | -34.2236
23.4526 | -35.4786 | -34.2299
23.4587 | -35.4791 | -34.2361
23.4648 | -35.4795 | -34.242
23.4709 | -35.4799 | -34.2481
23.477 | -35.4804 | -34.2543
23.4831 | -35.4808 | -34.2604
23.4892 | -35.4813 | -34.2665
23.4953 | -35.4817 | -34.2726
23.5014 | -35.4821 | -34.2786
23.5075 | -35.4826 | -34.2848
23.5136 | -35.483 | -34.2908
23.5197 | -35.4835 | -34.2967
23.5258 | -35.4839 | -34.3028
23.5319 | -35.4844 | -34.3086
23.538 | -35.4848 | -34.3147
23.5441 | -35.4852 | -34.3209
23.5502 | -35.4857 | -34.327
23.5563 | -35.4861 | -34.3332
23.5624 | -35.4865 | -34.3394
23.5686 | -35.487 | -34.3455
23.5747 | -35.4874 | -34.3514
23.5808 | -35.4879 | -34.3575
23.5869 | -35.4883 | -34.3634
23.593 | -35.4888 | -34.3696
23.5991 | -35.4892 | -34.3758
23.6052 | -35.4896 | -34.3817
23.6113 | -35.4901 | -34.3879
23.6174 | -35.4905 | -34.394
23.6235 | -35.491 | -34.4001
23.6296 | -35.4914 | -34.4064
23.6357 | -35.4918 | -34.4124
23.6418 | -35.4923 | -34.4183
23.6479 | -35.4927 | -34.4242
23.654 | -35.4932 | -34.4303
23.6601 | -35.4936 | -34.4364
23.6662 | -35.4941 | -34.4424
23.6723 | -35.4945 | -34.4483
23.6784 | -35.4949 | -34.4543
23.6845 | -35.4954 | -34.4604
23.6906 | -35.4958 | -34.4665
23.6967 | -35.4962 | -34.4726
23.7028 | -35.4967 | -34.4786
23.7089 | -35.4971 | -34.4845
23.7151 | -35.4975 | -34.4905
23.7212 | -35.498 | -34.4967
23.7273 | -35.4984 | -34.5028
23.7334 | -35.4988 | -34.5088
23.7395 | -35.4993 | -34.515
23.7456 | -35.4997 | -34.5208
23.7517 | -35.5001 | -34.5269
23.7578 | -35.5005 | -34.5329
23.7639 | -35.501 | -34.539
23.77 | -35.5014 | -34.5451
23.7761 | -35.5019 | -34.5513
23.7822 | -35.5023 | -34.5571
23.7883 | -35.5027 | -34.563
23.7944 | -35.5032 | -34.569
23.8005 | -35.5036 | -34.5751
23.8066 | -35.5041 | -34.5812
23.8127 | -35.5045 | -34.5871
23.8188 | -35.5049 | -34.5933
23.8249 | -35.5053 | -34.5994
23.831 | -35.5058 | -34.6054
23.8371 | -35.5062 | -34.6116
23.8432 | -35.5066 | -34.6176
23.8493 | -35.507 | -34.6235
23.8555 | -35.5074 | -34.6295
23.8616 | -35.5079 | -34.6355
23.8677 | -35.5083 | -34.6415
23.8738 | -35.5087 | -34.6475
23.8799 | -35.5091 | -34.6535
23.886 | -35.5096 | -34.6593
23.8921 | -35.51 | -34.6652
23.8982 | -35.5104 | -34.6713
23.9043 | -35.5109 | -34.6774
23.9104 | -35.5113 | -34.6833
23.9165 | -35.5117 | -34.6893
23.9226 | -35.5121 | -34.6951
23.9287 | -35.5126 | -34.7012
23.9348 | -35.513 | -34.7071
23.9409 | -35.5134 | -34.7131
23.947 | -35.5139 | -34.7193
23.9531 | -35.5143 | -34.7252
23.9592 | -35.5147 | -34.7312
23.9653 | -35.5152 | -34.7373
23.9714 | -35.5156 | -34.7434
23.9775 | -35.516 | -34.7494
23.9836 | -35.5165 | -34.7554
23.9897 | -35.5169 | -34.7613
23.9958 | -35.5173 | -34.7673
24.002 | -35.5178 | -34.7733
24.0081 | -35.5182 | -34.7795
24.0142 | -35.5186 | -34.7855
24.0203 | -35.519 | -34.7916
24.0264 | -35.5195 | -34.7973
24.0325 | -35.5199 | -34.8034
24.0386 | -35.5203 | -34.8095
24.0447 | -35.5207 | -34.8154
24.0508 | -35.5212 | -34.8213
24.0569 | -35.5216 | -34.8273
24.063 | -35.522 | -34.8332
24.0691 | -35.5225 | -34.8392
24.0752 | -35.5229 | -34.8453
24.0813 | -35.5233 | -34.8511
24.0874 | -35.5238 | -34.8571
24.0935 | -35.5242 | -34.8631
24.0996 | -35.5246 | -34.8688
24.1057 | -35.525 | -34.8748
24.1118 | -35.5255 | -34.8809
24.1179 | -35.5259 | -34.887
24.124 | -35.5263 | -34.8929
24.1301 | -35.5267 | -34.8988
24.1362 | -35.5272 | -34.9048
24.1424 | -35.5276 | -34.9109
24.1485 | -35.528 | -34.9172
24.1546 | -35.5285 | -34.9231
24.1607 | -35.5289 | -34.9292
24.1668 | -35.5293 | -34.935
24.1729 | -35.5297 | -34.9411
24.179 | -35.5301 | -34.9469
24.1851 | -35.5305 | -34.9528
24.1912 | -35.5309 | -34.9587
24.1973 | -35.5313 | -34.9647
24.2034 | -35.5318 | -34.9706
24.2095 | -35.5322 | -34.9767
24.2156 | -35.5326 | -34.9828
24.2217 | -35.533 | -34.9888
24.2278 | -35.5335 | -34.9948
24.2339 | -35.5339 | -35.0007
24.24 | -35.5343 | -35.0065
24.2461 | -35.5347 | -35.0123
24.2522 | -35.5352 | -35.0184
24.2583 | -35.5356 | -35.0243
24.2644 | -35.536 | -35.0305
24.2705 | -35.5365 | -35.0364
24.2766 | -35.5369 | -35.0424
24.2827 | -35.5373 | -35.0484
24.2889 | -35.5377 | -35.0545
24.295 | -35.5382 | -35.0603
24.3011 | -35.5386 | -35.0663
24.3072 | -35.539 | -35.0721
24.3133 | -35.5394 | -35.0781
24.3194 | -35.5398 | -35.0841
24.3255 | -35.5402 | -35.0902
24.3316 | -35.5406 | -35.0962
24.3377 | -35.541 | -35.1021
24.3438 | -35.5414 | -35.108
24.3499 | -35.5419 | -35.114
24.356 | -35.5423 | -35.1199
24.3621 | -35.5427 | -35.1259
24.3682 | -35.5431 | -35.1319
24.3743 | -35.5435 | -35.1379
24.3804 | -35.5439 | -35.1436
24.3865 | -35.5443 | -35.1496
24.3926 | -35.5448 | -35.1556
24.3987 | -35.5452 | -35.1615
24.4048 | -35.5456 | -35.1674
24.4109 | -35.546 | -35.1736
24.417 | -35.5464 | -35.1794
24.4231 | -35.5468 | -35.1854
24.4293 | -35.5473 | -35.1912
24.4354 | -35.5477 | -35.1973
24.4415 | -35.5481 | -35.2033
24.4476 | -35.5485 | -35.2091
24.4537 | -35.5489 | -35.2149
24.4598 | -35.5493 | -35.2208
24.4659 | -35.5497 | -35.2268
24.472 | -35.5501 | -35.2329
24.4781 | -35.5505 | -35.2388
24.4842 | -35.5509 | -35.2448
24.4903 | -35.5513 | -35.2507
24.4964 | -35.5518 | -35.2567
24.5025 | -35.5522 | -35.2626
24.5086 | -35.5526 | -35.2686
24.5147 | -35.553 | -35.2745
24.5208 | -35.5534 | -35.2801
24.5269 | -35.5538 | -35.286
24.533 | -35.5543 | -35.2918
24.5391 | -35.5547 | -35.2978
24.5452 | -35.5551 | -35.3037
24.5513 | -35.5555 | -35.3095
24.5574 | -35.5559 | -35.3153
24.5635 | -35.5564 | -35.3212
24.5696 | -35.5568 | -35.3271
24.5758 | -35.5572 | -35.3329
24.5819 | -35.5576 | -35.3388
24.588 | -35.558 | -35.3446
24.5941 | -35.5584 | -35.3505
24.6002 | -35.5588 | -35.3564
24.6063 | -35.5593 | -35.3624
24.6124 | -35.5597 | -35.3683
24.6185 | -35.5601 | -35.3741
24.6246 | -35.5605 | -35.38
24.6307 | -35.5609 | -35.3861
24.6368 | -35.5613 | -35.392
24.6429 | -35.5618 | -35.3978
24.649 | -35.5622 | -35.4036
24.6551 | -35.5626 | -35.4095
24.6612 | -35.563 | -35.4152
24.6673 | -35.5634 | -35.421
24.6734 | -35.5638 | -35.4269
24.6795 | -35.5642 | -35.4328
24.6856 | -35.5646 | -35.4388
24.6917 | -35.565 | -35.4448
24.6978 | -35.5654 | -35.4505
24.7039 | -35.5659 | -35.4564
24.71 | -35.5663 | -35.4623
24.7162 | -35.5667 | -35.4683
24.7223 | -35.5671 | -35.4741
24.7284 | -35.5675 | -35.4799
24.7345 | -35.5679 | -35.4858
24.7406 | -35.5683 | -35.4917
24.7467 | -35.5687 | -35.4977
24.7528 | -35.5691 | -35.5035
24.7589 | -35.5695 | -35.5093
24.765 | -35.5699 | -35.515
24.7711 | -35.5703 | -35.5209
24.7772 | -35.5707 | -35.5267
24.7833 | -35.5711 | -35.5325
24.7894 | -35.5715 | -35.5384
24.7955 | -35.5719 | -35.5443
24.8016 | -35.5723 | -35.55
24.8077 | -35.5727 | -35.5558
24.8138 | -35.5731 | -35.5617
24.8199 | -35.5735 | -35.5676
24.826 | -35.5739 | -35.5734
24.8321 | -35.5743 | -35.5792
24.8382 | -35.5747 | -35.5849
24.8443 | -35.5751 | -35.5907
24.8504 | -35.5755 | -35.5964
24.8565 | -35.5759 | -35.6024
24.8627 | -35.5764 | -35.6081
24.8688 | -35.5767 | -35.614
24.8749 | -35.5772 | -35.6199
24.881 | -35.5776 | -35.6258
24.8871 | -35.578 | -35.6317
24.8932 | -35.5784 | -35.6376
24.8993 | -35.5788 | -35.6433
24.9054 | -35.5792 | -35.6492
24.9115 | -35.5796 | -35.655
24.9176 | -35.5801 | -35.6609
24.9237 | -35.5805 | -35.6666
24.9298 | -35.5809 | -35.6725
24.9359 | -35.5813 | -35.6784
24.942 | -35.5817 | -35.6841
24.9481 | -35.5821 | -35.6901
24.9542 | -35.5825 | -35.696
24.9603 | -35.5829 | -35.7018
24.9664 | -35.5833 | -35.7077
24.9725 | -35.5837 | -35.7135
24.9786 | -35.584 | -35.7192
24.9847 | -35.5844 | -35.725
24.9908 | -35.5848 | -35.7309
24.9969 | -35.5852 | -35.7366
25.0031 | -35.5856 | -35.7425
25.0092 | -35.586 | -35.7483
25.0153 | -35.5864 | -35.754
25.0214 | -35.5868 | -35.7599
25.0275 | -35.5872 | -35.7657
25.0336 | -35.5876 | -35.7715
25.0397 | -35.588 | -35.7775
25.0458 | -35.5884 | -35.7832
25.0519 | -35.5889 | -35.789
25.058 | -35.5892 | -35.7949
25.0641 | -35.5896 | -35.8008
25.0702 | -35.59 | -35.8067
25.0763 | -35.5904 | -35.8126
25.0824 | -35.5908 | -35.8184
25.0885 | -35.5912 | -35.8242
25.0946 | -35.5916 | -35.8301
25.1007 | -35.592 | -35.8358
25.1068 | -35.5924 | -35.8417
25.1129 | -35.5928 | -35.8475
25.119 | -35.5932 | -35.8531
25.1251 | -35.5936 | -35.859
25.1312 | -35.594 | -35.8649
25.1373 | -35.5944 | -35.8708
25.1435 | -35.5948 | -35.8767
25.1496 | -35.5951 | -35.8824
25.1557 | -35.5955 | -35.8883
25.1618 | -35.5959 | -35.8942
25.1679 | -35.5963 | -35.9001
25.174 | -35.5967 | -35.9059
25.1801 | -35.5971 | -35.9117
25.1862 | -35.5975 | -35.9175
25.1923 | -35.5978 | -35.9232
25.1984 | -35.5982 | -35.9291
25.2045 | -35.5986 | -35.9348
25.2106 | -35.599 | -35.9408
25.2167 | -35.5994 | -35.9467
25.2228 | -35.5998 | -35.9524
25.2289 | -35.6002 | -35.9582
25.235 | -35.6006 | -35.9641
25.2411 | -35.601 | -35.9699
25.2472 | -35.6014 | -35.9757
25.2533 | -35.6018 | -35.9817
25.2594 | -35.6022 | -35.9873
25.2655 | -35.6026 | -35.9932
25.2716 | -35.603 | -35.999
25.2777 | -35.6034 | -36.0049
25.2838 | -35.6038 | -36.0108
25.29 | -35.6042 | -36.0165
25.2961 | -35.6046 | -36.0221
25.3022 | -35.605 | -36.0279
25.3083 | -35.6054 | -36.0338
25.3144 | -35.6058 | -36.0394
25.3205 | -35.6062 | -36.0454
25.3266 | -35.6065 | -36.0511
25.3327 | -35.6069 | -36.0571
25.3388 | -35.6073 | -36.063
25.3449 | -35.6077 | -36.0689
25.351 | -35.6081 | -36.0747
25.3571 | -35.6085 | -36.0805
25.3632 | -35.6089 | -36.086
25.3693 | -35.6093 | -36.092
25.3754 | -35.6097 | -36.0978
25.3815 | -35.6101 | -36.1037
25.3876 | -35.6105 | -36.1095
25.3937 | -35.6109 | -36.1152
25.3998 | -35.6112 | -36.1207
25.4059 | -35.6116 | -36.1266
25.412 | -35.612 | -36.1324
25.4181 | -35.6124 | -36.1382
25.4242 | -35.6128 | -36.144
25.4304 | -35.6131 | -36.1497
25.4365 | -35.6135 | -36.1553
25.4426 | -35.6139 | -36.161
25.4487 | -35.6143 | -36.1669
25.4548 | -35.6147 | -36.1725
25.4609 | -35.6151 | -36.1783
25.467 | -35.6154 | -36.1838
25.4731 | -35.6158 | -36.1896
25.4792 | -35.6162 | -36.1953
25.4853 | -35.6166 | -36.2012
25.4914 | -35.617 | -36.207
25.4975 | -35.6174 | -36.2127
25.5036 | -35.6178 | -36.2185
25.5097 | -35.6182 | -36.2244
25.5158 | -35.6185 | -36.2303
25.5219 | -35.6189 | -36.2362
25.528 | -35.6193 | -36.242
25.5341 | -35.6197 | -36.2478
25.5402 | -35.6201 | -36.2534
25.5463 | -35.6205 | -36.2592
25.5524 | -35.6209 | -36.265
25.5585 | -35.6213 | -36.2707
25.5646 | -35.6216 | -36.2766
25.5707 | -35.622 | -36.2825
25.5769 | -35.6224 | -36.2882
25.583 | -35.6228 | -36.294
25.5891 | -35.6232 | -36.2998
25.5952 | -35.6236 | -36.3055
25.6013 | -35.624 | -36.3112
25.6074 | -35.6244 | -36.3167
25.6135 | -35.6247 | -36.3224
25.6196 | -35.6251 | -36.3284
25.6257 | -35.6255 | -36.3341
25.6318 | -35.6259 | -36.3398
25.6379 | -35.6263 | -36.3456
25.644 | -35.6266 | -36.3512
25.6501 | -35.627 | -36.3569
25.6562 | -35.6274 | -36.3627
25.6623 | -35.6277 | -36.3683
25.6684 | -35.6281 | -36.3742
25.6745 | -35.6285 | -36.38
25.6806 | -35.6289 | -36.3855
25.6867 | -35.6293 | -36.3912
25.6928 | -35.6297 | -36.3969
25.6989 | -35.63 | -36.4027
25.705 | -35.6304 | -36.4086
25.7111 | -35.6308 | -36.4144
25.7173 | -35.6312 | -36.4199
25.7234 | -35.6316 | -36.4256
25.7295 | -35.632 | -36.4314
25.7356 | -35.6323 | -36.4373
25.7417 | -35.6327 | -36.4429
25.7478 | -35.6331 | -36.4487
25.7539 | -35.6335 | -36.4544
25.76 | -35.6338 | -36.4602
25.7661 | -35.6342 | -36.4658
25.7722 | -35.6346 | -36.4716
25.7783 | -35.6349 | -36.4773
25.7844 | -35.6353 | -36.4827
25.7905 | -35.6357 | -36.4885
25.7966 | -35.6361 | -36.4941
25.8027 | -35.6364 | -36.4998
25.8088 | -35.6368 | -36.5057
25.8149 | -35.6371 | -36.5114
25.821 | -35.6375 | -36.5171
25.8271 | -35.6379 | -36.5228
25.8332 | -35.6383 | -36.5284
25.8393 | -35.6386 | -36.534
25.8454 | -35.639 | -36.5397
25.8515 | -35.6394 | -36.5455
25.8576 | -35.6397 | -36.551
25.8638 | -35.6401 | -36.5567
25.8699 | -35.6405 | -36.5623
25.876 | -35.6409 | -36.568
25.8821 | -35.6412 | -36.5738
25.8882 | -35.6416 | -36.5792
25.8943 | -35.642 | -36.585
25.9004 | -35.6423 | -36.5908
25.9065 | -35.6427 | -36.5966
25.9126 | -35.6431 | -36.6022
25.9187 | -35.6434 | -36.608
25.9248 | -35.6438 | -36.6134
25.9309 | -35.6442 | -36.619
25.937 | -35.6446 | -36.6247
25.9431 | -35.6449 | -36.6304
25.9492 | -35.6453 | -36.6361
25.9553 | -35.6457 | -36.642
25.9614 | -35.6461 | -36.6476
25.9675 | -35.6464 | -36.6533
25.9736 | -35.6468 | -36.659
25.9797 | -35.6472 | -36.6647
25.9858 | -35.6475 | -36.6705
25.9919 | -35.6479 | -36.6762
25.998 | -35.6483 | -36.6818
26.0042 | -35.6487 | -36.6874
26.0103 | -35.649 | -36.6931
26.0164 | -35.6494 | -36.6988
26.0225 | -35.6498 | -36.7043
26.0286 | -35.6501 | -36.7099
26.0347 | -35.6505 | -36.7153
26.0408 | -35.6509 | -36.721
26.0469 | -35.6512 | -36.7264
26.053 | -35.6516 | -36.7321
26.0591 | -35.652 | -36.7377
26.0652 | -35.6524 | -36.7433
26.0713 | -35.6527 | -36.7489
26.0774 | -35.6531 | -36.7545
26.0835 | -35.6535 | -36.7603
26.0896 | -35.6538 | -36.7661
26.0957 | -35.6542 | -36.7718
26.1018 | -35.6546 | -36.7774
26.1079 | -35.6549 | -36.783
26.114 | -35.6553 | -36.7886
26.1201 | -35.6557 | -36.7942
26.1262 | -35.6561 | -36.7999
26.1323 | -35.6564 | -36.8055
26.1384 | -35.6568 | -36.8109
26.1445 | -35.6571 | -36.8167
26.1507 | -35.6575 | -36.8224
26.1568 | -35.6579 | -36.8281
26.1629 | -35.6582 | -36.8339
26.169 | -35.6586 | -36.8394
26.1751 | -35.6589 | -36.8448
26.1812 | -35.6593 | -36.8505
26.1873 | -35.6597 | -36.8562
26.1934 | -35.66 | -36.8617
26.1995 | -35.6604 | -36.8673
26.2056 | -35.6608 | -36.873
26.2117 | -35.6611 | -36.8786
26.2178 | -35.6615 | -36.8841
26.2239 | -35.6619 | -36.8898
26.23 | -35.6622 | -36.8955
26.2361 | -35.6626 | -36.9014
26.2422 | -35.663 | -36.9069
26.2483 | -35.6633 | -36.9125
26.2544 | -35.6637 | -36.918
26.2605 | -35.6641 | -36.9235
26.2666 | -35.6645 | -36.9291
26.2727 | -35.6648 | -36.9347
26.2788 | -35.6652 | -36.94
26.2849 | -35.6655 | -36.9456
26.2911 | -35.6659 | -36.9513
26.2972 | -35.6663 | -36.957
26.3033 | -35.6667 | -36.9627
26.3094 | -35.667 | -36.9684
26.3155 | -35.6673 | -36.9737
26.3216 | -35.6677 | -36.9794
26.3277 | -35.668 | -36.9849
26.3338 | -35.6684 | -36.9905
26.3399 | -35.6688 | -36.9961
26.346 | -35.6691 | -37.0015
26.3521 | -35.6694 | -37.007
26.3582 | -35.6698 | -37.0127
26.3643 | -35.6702 | -37.0183
26.3704 | -35.6706 | -37.024
26.3765 | -35.6709 | -37.0295
26.3826 | -35.6713 | -37.0348
26.3887 | -35.6717 | -37.0405
26.3948 | -35.672 | -37.0461
26.4009 | -35.6724 | -37.0517
26.407 | -35.6727 | -37.0573
26.4131 | -35.6731 | -37.0631
26.4192 | -35.6734 | -37.0685
26.4253 | -35.6738 | -37.0742
26.4314 | -35.6742 | -37.0797
26.4376 | -35.6745 | -37.0853
26.4437 | -35.6749 | -37.0909
26.4498 | -35.6752 | -37.0966
26.4559 | -35.6756 | -37.1021
26.462 | -35.6759 | -37.1076
26.4681 | -35.6763 | -37.1132
26.4742 | -35.6767 | -37.1188
26.4803 | -35.677 | -37.1245
26.4864 | -35.6773 | -37.1299
26.4925 | -35.6777 | -37.1354
26.4986 | -35.6781 | -37.1411
26.5047 | -35.6784 | -37.1467
26.5108 | -35.6788 | -37.1523
26.5169 | -35.6791 | -37.1578
26.523 | -35.6795 | -37.1632
26.5291 | -35.6798 | -37.1688
26.5352 | -35.6801 | -37.1742
26.5413 | -35.6805 | -37.1797
26.5474 | -35.6808 | -37.1853
26.5535 | -35.6812 | -37.1909
26.5596 | -35.6816 | -37.1963
26.5657 | -35.6819 | -37.2019
26.5718 | -35.6823 | -37.2072
26.578 | -35.6827 | -37.2128
26.5841 | -35.683 | -37.2183
26.5902 | -35.6834 | -37.224
26.5963 | -35.6837 | -37.2293
26.6024 | -35.6841 | -37.235
26.6085 | -35.6844 | -37.2406
26.6146 | -35.6848 | -37.2462
26.6207 | -35.6852 | -37.2519
26.6268 | -35.6855 | -37.2575
26.6329 | -35.6858 | -37.2628
26.639 | -35.6862 | -37.2683
26.6451 | -35.6866 | -37.2739
26.6512 | -35.6869 | -37.2793
26.6573 | -35.6873 | -37.2849
26.6634 | -35.6876 | -37.2904
26.6695 | -35.688 | -37.296
26.6756 | -35.6883 | -37.3017
26.6817 | -35.6887 | -37.3072
26.6878 | -35.689 | -37.3126
26.6939 | -35.6894 | -37.3182
26.7 | -35.6897 | -37.3236
26.7061 | -35.6901 | -37.329
26.7122 | -35.6905 | -37.3345
26.7183 | -35.6908 | -37.3401
26.7245 | -35.6912 | -37.3457
26.7306 | -35.6915 | -37.3512
26.7367 | -35.6918 | -37.3567
26.7428 | -35.6922 | -37.3622
26.7489 | -35.6926 | -37.3677
26.755 | -35.6929 | -37.3732
26.7611 | -35.6932 | -37.3788
26.7672 | -35.6936 | -37.3846
26.7733 | -35.6939 | -37.3899
26.7794 | -35.6943 | -37.3954
26.7855 | -35.6946 | -37.4009
26.7916 | -35.695 | -37.4065
26.7977 | -35.6953 | -37.4121
26.8038 | -35.6957 | -37.4175
26.8099 | -35.696 | -37.4231
26.816 | -35.6964 | -37.4286
26.8221 | -35.6967 | -37.434
26.8282 | -35.6971 | -37.4394
26.8343 | -35.6974 | -37.445
26.8404 | -35.6978 | -37.4503
26.8465 | -35.6981 | -37.4558
26.8526 | -35.6984 | -37.4613
26.8587 | -35.6988 | -37.4668
26.8649 | -35.6991 | -37.4724
26.871 | -35.6995 | -37.4779
26.8771 | -35.6998 | -37.4832
26.8832 | -35.7001 | -37.4887
26.8893 | -35.7005 | -37.4942
26.8954 | -35.7009 | -37.4998
26.9015 | -35.7012 | -37.5052
26.9076 | -35.7015 | -37.5107
26.9137 | -35.7019 | -37.516
26.9198 | -35.7022 | -37.5217
26.9259 | -35.7025 | -37.5272
26.932 | -35.7029 | -37.5327
26.9381 | -35.7032 | -37.5383
26.9442 | -35.7035 | -37.5437
26.9503 | -35.7039 | -37.5491
26.9564 | -35.7042 | -37.5544
26.9625 | -35.7046 | -37.5601
26.9686 | -35.7049 | -37.5657
26.9747 | -35.7053 | -37.5712
26.9808 | -35.7056 | -37.5765
26.9869 | -35.706 | -37.582
26.993 | -35.7063 | -37.5874
26.9991 | -35.7067 | -37.5928
27.0052 | -35.707 | -37.5983
27.0114 | -35.7074 | -37.6037
27.0175 | -35.7077 | -37.609
27.0236 | -35.708 | -37.6146
27.0297 | -35.7084 | -37.6201
27.0358 | -35.7087 | -37.6257
27.0419 | -35.7091 | -37.6311
27.048 | -35.7094 | -37.6367
27.0541 | -35.7098 | -37.642
27.0602 | -35.7101 | -37.6474
27.0663 | -35.7104 | -37.653
27.0724 | -35.7108 | -37.6585
27.0785 | -35.7111 | -37.6641
27.0846 | -35.7115 | -37.6694
27.0907 | -35.7119 | -37.675
27.0968 | -35.7122 | -37.6805
27.1029 | -35.7125 | -37.6858
27.109 | -35.7128 | -37.6912
27.1151 | -35.7132 | -37.6966
27.1212 | -35.7135 | -37.702
27.1273 | -35.7139 | -37.7074
27.1334 | -35.7142 | -37.7128
27.1395 | -35.7145 | -37.7184
27.1456 | -35.7149 | -37.7238
27.1518 | -35.7153 | -37.7292
27.1579 | -35.7156 | -37.7345
27.164 | -35.716 | -37.74
27.1701 | -35.7163 | -37.7454
27.1762 | -35.7166 | -37.7508
27.1823 | -35.717 | -37.7563
27.1884 | -35.7173 | -37.7619
27.1945 | -35.7177 | -37.7671
27.2006 | -35.718 | -37.7725
27.2067 | -35.7184 | -37.7781
27.2128 | -35.7187 | -37.7837
27.2189 | -35.7191 | -37.7891
27.225 | -35.7194 | -37.7946
27.2311 | -35.7197 | -37.7999
27.2372 | -35.7201 | -37.8054
27.2433 | -35.7205 | -37.8107
27.2494 | -35.7208 | -37.8163
27.2555 | -35.7211 | -37.8219
27.2616 | -35.7215 | -37.8271
27.2677 | -35.7218 | -37.8324
27.2738 | -35.7222 | -37.8376
27.2799 | -35.7225 | -37.8433
27.286 | -35.7228 | -37.8487
27.2921 | -35.7232 | -37.8541
27.2983 | -35.7235 | -37.8592
27.3044 | -35.7239 | -37.8646
27.3105 | -35.7242 | -37.8699
27.3166 | -35.7246 | -37.8752
27.3227 | -35.7249 | -37.8806
27.3288 | -35.7253 | -37.886
27.3349 | -35.7256 | -37.8911
27.341 | -35.7259 | -37.8964
27.3471 | -35.7262 | -37.9017
27.3532 | -35.7266 | -37.9069
27.3593 | -35.7269 | -37.9126
27.3654 | -35.7273 | -37.918
27.3715 | -35.7276 | -37.9231
27.3776 | -35.728 | -37.9286
27.3837 | -35.7283 | -37.9338
27.3898 | -35.7287 | -37.9392
27.3959 | -35.729 | -37.9447
27.402 | -35.7293 | -37.9497
27.4081 | -35.7297 | -37.955
27.4142 | -35.73 | -37.9604
27.4203 | -35.7303 | -37.9657
27.4264 | -35.7307 | -37.9712
27.4325 | -35.731 | -37.9765
27.4387 | -35.7313 | -37.9818
27.4448 | -35.7317 | -37.9872
27.4509 | -35.732 | -37.9925
27.457 | -35.7324 | -37.9978
27.4631 | -35.7327 | -38.003
27.4692 | -35.733 | -38.0084
27.4753 | -35.7334 | -38.0137
27.4814 | -35.7337 | -38.019
27.4875 | -35.7341 | -38.0246
27.4936 | -35.7344 | -38.0299
27.4997 | -35.7347 | -38.0353
27.5058 | -35.7351 | -38.0407
27.5119 | -35.7354 | -38.0458
27.518 | -35.7357 | -38.0511
27.5241 | -35.7361 | -38.0564
27.5302 | -35.7364 | -38.0619
27.5363 | -35.7367 | -38.0672
27.5424 | -35.737 | -38.0723
27.5485 | -35.7374 | -38.0778
27.5546 | -35.7377 | -38.0832
27.5607 | -35.7381 | -38.0886
27.5668 | -35.7384 | -38.094
27.5729 | -35.7388 | -38.0993
27.5791 | -35.7391 | -38.1044
27.5852 | -35.7394 | -38.1096
27.5913 | -35.7398 | -38.1149
27.5974 | -35.7401 | -38.1203
27.6035 | -35.7405 | -38.1258
27.6096 | -35.7408 | -38.1312
27.6157 | -35.7411 | -38.1362
27.6218 | -35.7415 | -38.1416
27.6279 | -35.7418 | -38.1469
27.634 | -35.7422 | -38.1522
27.6401 | -35.7425 | -38.1576
27.6462 | -35.7428 | -38.1629
27.6523 | -35.7431 | -38.1681
27.6584 | -35.7435 | -38.1733
27.6645 | -35.7439 | -38.1786
27.6706 | -35.7442 | -38.184
27.6767 | -35.7445 | -38.1893
27.6828 | -35.7448 | -38.1944
27.6889 | -35.7451 | -38.1996
27.695 | -35.7455 | -38.2049
27.7011 | -35.7458 | -38.2102
27.7072 | -35.7461 | -38.2154
27.7133 | -35.7465 | -38.2209
27.7194 | -35.7468 | -38.226
27.7256 | -35.7471 | -38.2313
27.7317 | -35.7475 | -38.2367
27.7378 | -35.7478 | -38.2421
27.7439 | -35.7481 | -38.2475
27.75 | -35.7485 | -38.2527
27.7561 | -35.7488 | -38.2577
27.7622 | -35.7491 | -38.2631
27.7683 | -35.7495 | -38.2683
27.7744 | -35.7498 | -38.2736
27.7805 | -35.7501 | -38.2788
27.7866 | -35.7504 | -38.2842
27.7927 | -35.7507 | -38.2893
27.7988 | -35.7511 | -38.2947
27.8049 | -35.7514 | -38.3001
27.811 | -35.7517 | -38.3053
27.8171 | -35.752 | -38.3107
27.8232 | -35.7524 | -38.3156
27.8293 | -35.7527 | -38.3207
27.8354 | -35.753 | -38.326
27.8415 | -35.7533 | -38.3313
27.8476 | -35.7536 | -38.3366
27.8537 | -35.7539 | -38.342
27.8598 | -35.7543 | -38.3472
27.866 | -35.7546 | -38.3525
27.8721 | -35.7549 | -38.3577
27.8782 | -35.7552 | -38.3631
27.8843 | -35.7556 | -38.3681
27.8904 | -35.756 | -38.3737
27.8965 | -35.7563 | -38.3791
27.9026 | -35.7566 | -38.3843
27.9087 | -35.757 | -38.3899
27.9148 | -35.7573 | -38.3949
27.9209 | -35.7575 | -38.4003
27.927 | -35.7579 | -38.4051
27.9331 | -35.7582 | -38.4106
27.9392 | -35.7586 | -38.4161
27.9453 | -35.7588 | -38.4214
27.9514 | -35.7591 | -38.4269
27.9575 | -35.7595 | -38.4324
27.9636 | -35.7598 | -38.4378
27.9697 | -35.7602 | -38.4432
27.9758 | -35.7605 | -38.4487
27.9819 | -35.7609 | -38.4542
27.988 | -35.7612 | -38.4591
27.9941 | -35.7615 | -38.4647
28.0002 | -35.7619 | -38.4693
28.0063 | -35.7622 | -38.4747
28.0125 | -35.7625 | -38.4803
28.0186 | -35.7629 | -38.4857
28.0247 | -35.7632 | -38.4912
28.0308 | -35.7635 | -38.4967
28.0369 | -35.7639 | -38.5019
28.043 | -35.7641 | -38.5071
28.0491 | -35.7644 | -38.5126
28.0552 | -35.7647 | -38.5179
28.0613 | -35.765 | -38.5233
28.0674 | -35.7653 | -38.5285
28.0735 | -35.7656 | -38.5338

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.538 | -43.4559
21.4908 | -35.5881 | -35.3957
16.4948 | -34.9633 | -27.9599
12.6602 | -34.6032 | -21.6545
9.7171 | -34.3787 | -16.7283
7.4582 | -34.2235 | -12.7312
5.7244 | -34.1066 | -9.4975
4.3936 | -34.0361 | -6.8832
3.3722 | -34.0133 | -4.8276
2.5883 | -34.0093 | -3.2658
1.9866 | -34.0085 | -2.0189
1.5248 | -33.998 | -1.0158
1.1703 | -34.0061 | -0.3651
1 | -34.0029 | -0.037583
0.89822 | -33.9962 | 0.23598
0.68941 | -33.9787 | 0.82169
0.52919 | -33.9796 | 1.2737
0.40612 | -33.9756 | 1.8878
0.31171 | -33.9692 | 2.3379
0.23928 | -33.9691 | 3.0671
0.18358 | -33.9813 | 3.4874
0.14091 | -33.9958 | 4.3666
0.10819 | -34.0145 | 5.5237
0.083017 | -34.0325 | 7.0014
0.063705 | -34.0622 | 8.8756
0.048923 | -34.1318 | 11.3753
0.037479 | -34.2391 | 14.5646
0.028801 | -34.4113 | 18.5818
0.022078 | -34.7171 | 23.8337
0.017023 | -35.1843 | 30.1761
0.013018 | -35.9696 | 37.855
0.010014 | -37.0814 | 45.777

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.00366 | -10.4511
0.00732 | -25.4054
0.01098 | -33.3083
0.014641 | -36.692
0.018301 | -39.1495
0.021961 | -42.0649
0.025621 | -45.9888
0.029281 | -46.9975
0.032941 | -47.0676
0.036602 | -48.9647
0.040262 | -50.3704
0.043922 | -51.3804
0.047582 | -52.0422
0.051242 | -53.8689
0.054902 | -55.05
0.058562 | -55.1765
0.062223 | -55.562
0.065883 | -56.9976
0.069543 | -57.6118
0.073203 | -58.8052
0.076863 | -58.8804
0.080523 | -57.5545
0.084184 | -57.1895
0.087844 | -59.991
0.091504 | -61.0151
0.095164 | -61.0435
0.098824 | -61.8108
0.10248 | -63.0922
0.10614 | -63.4941
0.1098 | -63.5349
0.11346 | -63.5703
0.11713 | -64.0015
0.12079 | -63.9357
0.12445 | -64.9321
0.1281 | -66.0393
0.13177 | -66.0139
0.13543 | -65.9411
0.13909 | -66.1812
0.14275 | -66.8453
0.14641 | -67.3494
0.15007 | -67.5504
0.15373 | -67.2226
0.15739 | -67.9779
0.16105 | -69.5087
0.16471 | -68.1299
0.16837 | -68.0092
0.17203 | -68.2283
0.17569 | -69.5585
0.17935 | -69.0617
0.18301 | -70.3529
0.18667 | -70.5166
0.19033 | -69.9059
0.19399 | -69.8058
0.19765 | -70.5375
0.20131 | -72.3086
0.20497 | -72.6929
0.20863 | -73.8278
0.21229 | -72.2733
0.21595 | -71.0128
0.21961 | -71.1953
0.22327 | -71.1656
0.22693 | -72.0783
0.23059 | -73.3155
0.23425 | -72.2988
0.23791 | -72.5483
0.24157 | -73.4966
0.24523 | -74.6414
0.24889 | -73.4722
0.25255 | -72.311
0.25621 | -73.4294
0.25987 | -74.5645
0.26353 | -74.0229
0.26719 | -74.1682
0.27085 | -74.0452
0.27451 | -75.5098
0.27817 | -75.5913
0.28183 | -74.3916
0.28549 | -74.222
0.28915 | -74.6021
0.29281 | -74.7927
0.29647 | -74.8595
0.30013 | -74.141
0.30379 | -74.8897
0.30745 | -74.7332
0.31111 | -74.496
0.31477 | -74.7249
0.31843 | -75.4546
0.32209 | -75.5158
0.32575 | -74.8792
0.32941 | -75.7203
0.33307 | -75.6694
0.33673 | -76.5314
0.34039 | -76.9568
0.34405 | -76.2113
0.34771 | -75.6376
0.35137 | -75.5765
0.35503 | -76.2741
0.35869 | -77.338
0.36235 | -76.9201
0.36601 | -76.545
0.36968 | -76.3873
0.37334 | -77.4076
0.377 | -76.9671
0.38066 | -77.4879
0.38432 | -77.1955
0.38798 | -76.2358
0.39164 | -76.1056
0.3953 | -76.982
0.39896 | -77.8762
0.40262 | -77.334
0.40628 | -77.235
0.40994 | -78.7347
0.4136 | -79.3024
0.41726 | -78.2226
0.42092 | -78.35
0.42458 | -78.039
0.42824 | -79.1828
0.4319 | -79.0497
0.43556 | -78.0319
0.43922 | -78.1481
0.44288 | -77.6326
0.44654 | -77.8913
0.4502 | -78.401
0.45386 | -78.7736
0.45752 | -78.1175
0.46118 | -79.1935
0.46484 | -79.9832
0.4685 | -79.3477
0.47216 | -78.9411
0.47582 | -79.4396
0.47948 | -79.5242
0.48314 | -79.208
0.4868 | -79.2459
0.49046 | -79.4999
0.49412 | -78.8656
0.49778 | -79.2295
0.50144 | -79.6849
0.5051 | -78.9783
0.50876 | -78.7678
0.51242 | -80.0061
0.51608 | -80.9014
0.51974 | -80.7167
0.5234 | -79.7959
0.52706 | -80.0347
0.53072 | -80.2653
0.53438 | -80.4657
0.53804 | -80.466
0.5417 | -80.4241
0.54536 | -80.0856
0.54902 | -79.9647
0.55268 | -80.4968
0.55634 | -81.1681
0.56 | -81.362
0.56366 | -81.6756
0.56732 | -82.1539
0.57098 | -82.2517
0.57464 | -81.8308
0.5783 | -81.6668
0.58196 | -81.6557
0.58562 | -81.6129
0.58928 | -81.7412
0.59295 | -82.1677
0.59661 | -82.4181
0.60027 | -82.1537
0.60393 | -81.8858
0.60759 | -82.084
0.61125 | -82.3095
0.61491 | -82.3457
0.61857 | -82.3318
0.62223 | -82.1607
0.62589 | -82.0753
0.62955 | -82.6635
0.63321 | -83.5465
0.63687 | -83.5947
0.64053 | -82.8373
0.64419 | -82.6501
0.64785 | -83.1484
0.65151 | -83.5047
0.65517 | -83.332
0.65883 | -82.9974
0.66249 | -82.5997
0.66615 | -82.277
0.66981 | -82.2872
0.67347 | -82.424
0.67713 | -82.5709
0.68079 | -82.508
0.68445 | -82.3312
0.68811 | -82.3835
0.69177 | -82.5505
0.69543 | -82.9426
0.69909 | -83.3755
0.70275 | -83.6536
0.70641 | -83.9408
0.71007 | -83.4569
0.71373 | -82.7152
0.71739 | -82.5145
0.72105 | -82.7509
0.72471 | -83.0373
0.72837 | -83.3778
0.73203 | -83.5134
0.73569 | -83.0656
0.73935 | -82.6953
0.74301 | -83.1825
0.74667 | -83.6454
0.75033 | -83.7495
0.75399 | -83.9608
0.75765 | -83.8662
0.76131 | -83.8492
0.76497 | -84.0071
0.76863 | -83.7932
0.77229 | -83.6862
0.77595 | -83.9145
0.77961 | -83.8851
0.78327 | -83.8899
0.78693 | -84.3002
0.79059 | -84.9884
0.79425 | -85.5378
0.79791 | -85.9978
0.80157 | -85.794
0.80523 | -84.8512
0.80889 | -84.2151
0.81255 | -84.4432
0.81621 | -85.2348
0.81988 | -85.6394
0.82354 | -84.9761
0.8272 | -84.2212
0.83086 | -84.5037
0.83452 | -85.4828
0.83818 | -85.6539
0.84184 | -85.7616
0.8455 | -86.053
0.84916 | -85.9704
0.85282 | -85.6965
0.85648 | -85.3082
0.86014 | -84.8731
0.8638 | -84.7357
0.86746 | -85.0041
0.87112 | -85.3226
0.87478 | -85.528
0.87844 | -85.5142
0.8821 | -85.5426
0.88576 | -85.8394
0.88942 | -86.0135
0.89308 | -86.2107
0.89674 | -86.5485
0.9004 | -86.8699
0.90406 | -86.9616
0.90772 | -86.5504
0.91138 | -86.1424
0.91504 | -86.1324
0.9187 | -86.7045
0.92236 | -86.8799
0.92602 | -86.1766
0.92968 | -85.6984
0.93334 | -85.7852
0.937 | -86.2417
0.94066 | -86.3396
0.94432 | -86.06
0.94798 | -86.3384
0.95164 | -87.0304
0.9553 | -87.1408
0.95896 | -86.8161
0.96262 | -86.4036
0.96628 | -85.9532
0.96994 | -85.7442
0.9736 | -85.8427
0.97726 | -85.8647
0.98092 | -86.0761
0.98458 | -86.7776
0.98824 | -87.6106
0.9919 | -87.6039
0.99556 | -86.8066
0.99922 | -86.517
1.0029 | -86.8777
1.0065 | -87.1111
1.0102 | -86.8562
1.0139 | -86.6194
1.0175 | -87.1758
1.0212 | -88.2763
1.0248 | -88.9415
1.0285 | -88.8951
1.0322 | -88.5564
1.0358 | -87.6434
1.0395 | -86.788
1.0431 | -86.87
1.0468 | -87.0675
1.0505 | -86.8662
1.0541 | -86.7712
1.0578 | -86.7557
1.0614 | -86.9773
1.0651 | -87.3825
1.0688 | -87.5889
1.0724 | -87.3221
1.0761 | -86.7865
1.0797 | -86.4658
1.0834 | -86.5596
1.0871 | -86.8666
1.0907 | -87.2809
1.0944 | -87.8621
1.098 | -88.4105
1.1017 | -88.6082
1.1054 | -88.4292
1.109 | -88.212
1.1127 | -88.3411
1.1163 | -88.7169
1.12 | -88.7818
1.1237 | -88.4375
1.1273 | -88.1236
1.131 | -88.0634
1.1346 | -88.1735
1.1383 | -88.2464
1.142 | -88.1494
1.1456 | -88.0432
1.1493 | -88.0979
1.1529 | -88.2285
1.1566 | -88.321
1.1603 | -88.4313
1.1639 | -88.5622
1.1676 | -88.5146
1.1712 | -88.2067
1.1749 | -87.952
1.1786 | -87.9954
1.1822 | -88.2089
1.1859 | -88.3972
1.1896 | -88.5443
1.1932 | -88.7421
1.1969 | -89.1256
1.2005 | -89.6161
1.2042 | -89.8746
1.2079 | -89.7636
1.2115 | -89.4361
1.2152 | -89.0199
1.2188 | -88.6359
1.2225 | -88.4596
1.2262 | -88.58
1.2298 | -88.8785
1.2335 | -89.096
1.2371 | -89.0402
1.2408 | -88.8859
1.2445 | -88.9274
1.2481 | -89.1094
1.2518 | -89.2114
1.2554 | -89.233
1.2591 | -89.1596
1.2628 | -88.9768
1.2664 | -88.893
1.2701 | -88.9819
1.2737 | -89.013
1.2774 | -88.8917
1.2811 | -88.8195
1.2847 | -88.8301
1.2884 | -88.8194
1.292 | -88.9148
1.2957 | -89.26
1.2994 | -89.6328
1.303 | -89.6946
1.3067 | -89.605
1.3103 | -89.7783
1.314 | -90.1276
1.3177 | -90.132
1.3213 | -89.678
1.325 | -89.2105
1.3286 | -89.0404
1.3323 | -89.1198
1.336 | -89.3252
1.3396 | -89.6629
1.3433 | -90.0759
1.3469 | -90.3384
1.3506 | -90.2545
1.3543 | -89.9908
1.3579 | -89.949
1.3616 | -90.2022
1.3652 | -90.4211
1.3689 | -90.4112
1.3726 | -90.4339
1.3762 | -90.5782
1.3799 | -90.565
1.3835 | -90.3631
1.3872 | -90.1831
1.3909 | -89.9825
1.3945 | -89.754
1.3982 | -89.7781
1.4018 | -90.1878
1.4055 | -90.658
1.4092 | -90.7569
1.4128 | -90.554
1.4165 | -90.5118
1.4201 | -90.8312
1.4238 | -91.337
1.4275 | -91.6347
1.4311 | -91.4622
1.4348 | -91.0922
1.4384 | -90.8658
1.4421 | -90.7839
1.4458 | -90.8088
1.4494 | -90.8582
1.4531 | -90.7362
1.4567 | -90.4875
1.4604 | -90.3864
1.4641 | -90.4748
1.4677 | -90.6235
1.4714 | -90.7942
1.475 | -90.9254
1.4787 | -90.902
1.4824 | -90.7275
1.486 | -90.5634
1.4897 | -90.553
1.4933 | -90.6242
1.497 | -90.6293
1.5007 | -90.6444
1.5043 | -90.755
1.508 | -90.895
1.5116 | -90.9809
1.5153 | -90.9675
1.519 | -90.8977
1.5226 | -90.9245
1.5263 | -91.0611
1.5299 | -91.1596
1.5336 | -91.0788
1.5373 | -90.8199
1.5409 | -90.5811
1.5446 | -90.5231
1.5482 | -90.6134
1.5519 | -90.8398
1.5556 | -91.1756
1.5592 | -91.3625
1.5629 | -91.2815
1.5665 | -91.1313
1.5702 | -91.0403
1.5739 | -91
1.5775 | -90.9984
1.5812 | -90.938
1.5848 | -90.8768
1.5885 | -90.861
1.5922 | -90.9352
1.5958 | -91.0881
1.5995 | -91.2153
1.6031 | -91.2192
1.6068 | -91.1276
1.6105 | -91.0589
1.6141 | -91.1057
1.6178 | -91.2635
1.6214 | -91.4631
1.6251 | -91.6254
1.6288 | -91.6712
1.6324 | -91.5966
1.6361 | -91.5087
1.6397 | -91.5092
1.6434 | -91.5928
1.6471 | -91.6176
1.6507 | -91.4877
1.6544 | -91.3229
1.6581 | -91.2874
1.6617 | -91.4416
1.6654 | -91.6941
1.669 | -91.8729
1.6727 | -91.8652
1.6764 | -91.7
1.68 | -91.5626
1.6837 | -91.5836
1.6873 | -91.7178
1.691 | -91.8539
1.6947 | -91.904
1.6983 | -91.8626
1.702 | -91.7212
1.7056 | -91.529
1.7093 | -91.3826
1.713 | -91.3518
1.7166 | -91.4581
1.7203 | -91.6307
1.7239 | -91.6991
1.7276 | -91.5744
1.7313 | -91.3925
1.7349 | -91.3886
1.7386 | -91.6232
1.7422 | -91.925
1.7459 | -92.0306
1.7496 | -91.8891
1.7532 | -91.6935
1.7569 | -91.5941
1.7605 | -91.5853
1.7642 | -91.5842
1.7679 | -91.5645
1.7715 | -91.6117
1.7752 | -91.7946
1.7788 | -92.0681
1.7825 | -92.254
1.7862 | -92.223
1.7898 | -92.0337
1.7935 | -91.8505
1.7971 | -91.7952
1.8008 | -91.8878
1.8045 | -92.0884
1.8081 | -92.3261
1.8118 | -92.5121
1.8154 | -92.6303
1.8191 | -92.7198
1.8228 | -92.7964
1.8264 | -92.8005
1.8301 | -92.6945
1.8337 | -92.5483
1.8374 | -92.4527
1.8411 | -92.4661
1.8447 | -92.6054
1.8484 | -92.8597
1.852 | -93.1652
1.8557 | -93.3853
1.8594 | -93.4142
1.863 | -93.2678
1.8667 | -93.0584
1.8703 | -92.8969
1.874 | -92.8034
1.8777 | -92.7158
1.8813 | -92.603
1.885 | -92.5506
1.8886 | -92.6528
1.8923 | -92.8234
1.896 | -92.8623
1.8996 | -92.7212
1.9033 | -92.5849
1.9069 | -92.5813
1.9106 | -92.6482
1.9143 | -92.6575
1.9179 | -92.5913
1.9216 | -92.527
1.9252 | -92.4935
1.9289 | -92.4808
1.9326 | -92.4765
1.9362 | -92.494
1.9399 | -92.5525
1.9435 | -92.6751
1.9472 | -92.8873
1.9509 | -93.1788
1.9545 | -93.4182
1.9582 | -93.4314
1.9618 | -93.1988
1.9655 | -92.9199
1.9692 | -92.7609
1.9728 | -92.7942
1.9765 | -93.0169
1.9801 | -93.3467
1.9838 | -93.6365
1.9875 | -93.7917
1.9911 | -93.8287
1.9948 | -93.8404
1.9984 | -93.9008
2.0021 | -93.9896
2.0058 | -93.9866
2.0094 | -93.8623
2.0131 | -93.73
2.0167 | -93.641
2.0204 | -93.5298
2.0241 | -93.3398
2.0277 | -93.1209
2.0314 | -92.9298
2.035 | -92.7887
2.0387 | -92.7332
2.0424 | -92.8101
2.046 | -93.0013
2.0497 | -93.2293
2.0533 | -93.4023
2.057 | -93.4681
2.0607 | -93.4391
2.0643 | -93.3743
2.068 | -93.3119
2.0716 | -93.2777
2.0753 | -93.2745
2.079 | -93.3018
2.0826 | -93.3168
2.0863 | -93.2814
2.0899 | -93.2248
2.0936 | -93.186
2.0973 | -93.2015
2.1009 | -93.2769
2.1046 | -93.3543
2.1082 | -93.3442
2.1119 | -93.2307
2.1156 | -93.1833
2.1192 | -93.1899
2.1229 | -93.2585
2.1265 | -93.3482
2.1302 | -93.4138
2.1339 | -93.4317
2.1375 | -93.4042
2.1412 | -93.3616
2.1449 | -93.3467
2.1485 | -93.3728
2.1522 | -93.4477
2.1558 | -93.5921
2.1595 | -93.8094
2.1632 | -94.0601
2.1668 | -94.2468
2.1705 | -94.2895
2.1741 | -94.1768
2.1778 | -93.9816
2.1815 | -93.7735
2.1851 | -93.5835
2.1888 | -93.4465
2.1924 | -93.3685
2.1961 | -93.3588
2.1998 | -93.4108
2.2034 | -93.5133
2.2071 | -93.6483
2.2107 | -93.7999
2.2144 | -93.9467
2.2181 | -94.0664
2.2217 | -94.1574
2.2254 | -94.2059
2.229 | -94.1759
2.2327 | -94.0988
2.2364 | -94.0449
2.24 | -94.0789
2.2437 | -94.1719
2.2473 | -94.2314
2.251 | -94.2165
2.2547 | -94.1626
2.2583 | -94.1355
2.262 | -94.1497
2.2656 | -94.2252
2.2693 | -94.3582
2.273 | -94.4961
2.2766 | -94.6052
2.2803 | -94.6756
2.2839 | -94.7286
2.2876 | -94.7475
2.2913 | -94.6759
2.2949 | -94.522
2.2986 | -94.3355
2.3022 | -94.2026
2.3059 | -94.1486
2.3096 | -94.1816
2.3132 | -94.2873
2.3169 | -94.3989
2.3205 | -94.4439
2.3242 | -94.3942
2.3279 | -94.2808
2.3315 | -94.1815
2.3352 | -94.1411
2.3388 | -94.176
2.3425 | -94.2422
2.3462 | -94.2655
2.3498 | -94.1701
2.3535 | -93.9521
2.3571 | -93.6847
2.3608 | -93.4682
2.3645 | -93.3743
2.3681 | -93.4374
2.3718 | -93.6411
2.3754 | -93.9453
2.3791 | -94.2769
2.3828 | -94.5707
2.3864 | -94.7375
2.3901 | -94.763
2.3937 | -94.6928
2.3974 | -94.6099
2.4011 | -94.5601
2.4047 | -94.567
2.4084 | -94.6276
2.412 | -94.722
2.4157 | -94.819
2.4194 | -94.896
2.423 | -94.9492
2.4267 | -94.9853
2.4303 | -94.9521
2.434 | -94.8328
2.4377 | -94.6565
2.4413 | -94.4989
2.445 | -94.4031
2.4486 | -94.3913
2.4523 | -94.438
2.456 | -94.4861
2.4596 | -94.4832
2.4633 | -94.4412
2.4669 | -94.4015
2.4706 | -94.3943
2.4743 | -94.4159
2.4779 | -94.4295
2.4816 | -94.4201
2.4852 | -94.4239
2.4889 | -94.4552
2.4926 | -94.5346
2.4962 | -94.6771
2.4999 | -94.8769
2.5035 | -95.0728
2.5072 | -95.2246
2.5109 | -95.2752
2.5145 | -95.2023
2.5182 | -95.0623
2.5218 | -94.9322
2.5255 | -94.884
2.5292 | -94.9489
2.5328 | -95.0706
2.5365 | -95.172
2.5401 | -95.2316
2.5438 | -95.2707
2.5475 | -95.292
2.5511 | -95.3087
2.5548 | -95.2975
2.5584 | -95.2376
2.5621 | -95.1199
2.5658 | -94.9389
2.5694 | -94.7373
2.5731 | -94.5702
2.5767 | -94.475
2.5804 | -94.4586
2.5841 | -94.5303
2.5877 | -94.6841
2.5914 | -94.883
2.5951 | -95.0563
2.5987 | -95.1372
2.6024 | -95.1255
2.606 | -95.0907
2.6097 | -95.0811
2.6134 | -95.1144
2.617 | -95.1745
2.6207 | -95.2317
2.6243 | -95.2295
2.628 | -95.1565
2.6317 | -95.0765
2.6353 | -95.0791
2.639 | -95.1959
2.6426 | -95.2572
2.6463 | -95.3038
2.65 | -95.3012
2.6536 | -95.2593
2.6573 | -95.22
2.6609 | -95.2036
2.6646 | -95.201
2.6683 | -95.1947
2.6719 | -95.1599
2.6756 | -95.1011
2.6792 | -95.0451
2.6829 | -95.0294
2.6866 | -95.0805
2.6902 | -95.167
2.6939 | -95.2596
2.6975 | -95.3444
2.7012 | -95.4091
2.7049 | -95.4651
2.7085 | -95.5383
2.7122 | -95.6484
2.7158 | -95.7753
2.7195 | -95.8788
2.7232 | -95.9288
2.7268 | -95.9395
2.7305 | -95.9404
2.7341 | -95.9467
2.7378 | -95.9929
2.7415 | -96.0892
2.7451 | -96.2178
2.7488 | -96.3437
2.7524 | -96.4319
2.7561 | -96.4411
2.7598 | -96.3693
2.7634 | -96.2399
2.7671 | -96.0859
2.7707 | -95.9348
2.7744 | -95.8118
2.7781 | -95.7163
2.7817 | -95.6424
2.7854 | -95.5954
2.789 | -95.5555
2.7927 | -95.523
2.7964 | -95.5018
2.8 | -95.4931
2.8037 | -95.5052
2.8073 | -95.5358
2.811 | -95.5978
2.8147 | -95.6925
2.8183 | -95.8169
2.822 | -95.9626
2.8256 | -96.103
2.8293 | -96.2109
2.833 | -96.2391
2.8366 | -96.1977
2.8403 | -96.0862
2.8439 | -95.9374
2.8476 | -95.7944
2.8513 | -95.6969
2.8549 | -95.6565
2.8586 | -95.6732
2.8622 | -95.7292
2.8659 | -95.8074
2.8696 | -95.9066
2.8732 | -96.0291
2.8769 | -96.1609
2.8805 | -96.2967
2.8842 | -96.3934
2.8879 | -96.4377
2.8915 | -96.3829
2.8952 | -96.2448
2.8988 | -96.0609
2.9025 | -95.8755
2.9062 | -95.7251
2.9098 | -95.6283
2.9135 | -95.5858
2.9171 | -95.5749
2.9208 | -95.6019
2.9245 | -95.6759
2.9281 | -95.7888
2.9318 | -95.9322
2.9354 | -96.0555
2.9391 | -96.1407
2.9428 | -96.1911
2.9464 | -96.2486
2.9501 | -96.3458
2.9537 | -96.4894
2.9574 | -96.6482
2.9611 | -96.7655
2.9647 | -96.8104
2.9684 | -96.7567
2.972 | -96.6377
2.9757 | -96.5194
2.9794 | -96.4364
2.983 | -96.4105
2.9867 | -96.4304
2.9903 | -96.4651
2.994 | -96.4937
2.9977 | -96.4995
3.0013 | -96.4893
3.005 | -96.4768
3.0086 | -96.462
3.0123 | -96.4409
3.016 | -96.4134
3.0196 | -96.3778
3.0233 | -96.3467
3.0269 | -96.3374
3.0306 | -96.3508
3.0343 | -96.3694
3.0379 | -96.3535
3.0416 | -96.3012
3.0452 | -96.2149
3.0489 | -96.1118
3.0526 | -96.0224
3.0562 | -95.977
3.0599 | -95.987
3.0635 | -96.0422
3.0672 | -96.1127
3.0709 | -96.1599
3.0745 | -96.1723
3.0782 | -96.1665
3.0819 | -96.1795
3.0855 | -96.2555
3.0892 | -96.3994
3.0928 | -96.6204
3.0965 | -96.8829
3.1002 | -97.1378
3.1038 | -97.3292
3.1075 | -97.4167
3.1111 | -97.4309
3.1148 | -97.4259
3.1185 | -97.4515
3.1221 | -97.5409
3.1258 | -97.676
3.1294 | -97.8283
3.1331 | -97.927
3.1368 | -97.9414
3.1404 | -97.8938
3.1441 | -97.8256
3.1477 | -97.764
3.1514 | -97.7066
3.1551 | -97.6265
3.1587 | -97.4921
3.1624 | -97.3044
3.166 | -97.0894
3.1697 | -96.9373
3.1734 | -96.824
3.177 | -96.7756
3.1807 | -96.7917
3.1843 | -96.8765
3.188 | -96.9962
3.1917 | -97.1193
3.1953 | -97.2175
3.199 | -97.2653
3.2026 | -97.264
3.2063 | -97.2329
3.21 | -97.1797
3.2136 | -97.1185
3.2173 | -97.0549
3.2209 | -96.99
3.2246 | -96.921
3.2283 | -96.8438
3.2319 | -96.7741
3.2356 | -96.7185
3.2392 | -96.6868
3.2429 | -96.7096
3.2466 | -96.7925
3.2502 | -96.9408
3.2539 | -97.1442
3.2575 | -97.3711
3.2612 | -97.5865
3.2649 | -97.7496
3.2685 | -97.8359
3.2722 | -97.843
3.2758 | -97.7863
3.2795 | -97.6873
3.2832 | -97.5763
3.2868 | -97.4797
3.2905 | -97.409
3.2941 | -97.365
3.2978 | -97.3241
3.3015 | -97.2945
3.3051 | -97.2863
3.3088 | -97.2874
3.3124 | -97.2942
3.3161 | -97.295
3.3198 | -97.2934
3.3234 | -97.2767
3.3271 | -97.2373
3.3307 | -97.2012
3.3344 | -97.1894
3.3381 | -97.1948
3.3417 | -97.2076
3.3454 | -97.2494
3.349 | -97.3031
3.3527 | -97.3544
3.3564 | -97.4002
3.36 | -97.4338
3.3637 | -97.4743
3.3673 | -97.5353
3.371 | -97.6339
3.3747 | -97.7443
3.3783 | -97.8592
3.382 | -97.9595
3.3856 | -98.0055
3.3893 | -97.9844
3.393 | -97.9059
3.3966 | -97.79
3.4003 | -97.6509
3.4039 | -97.5208
3.4076 | -97.421
3.4113 | -97.356
3.4149 | -97.3413
3.4186 | -97.3738
3.4222 | -97.4238
3.4259 | -97.4634
3.4296 | -97.4948
3.4332 | -97.4932
3.4369 | -97.4658
3.4405 | -97.4248
3.4442 | -97.3855
3.4479 | -97.36
3.4515 | -97.3512
3.4552 | -97.3677
3.4588 | -97.4038
3.4625 | -97.458
3.4662 | -97.5268
3.4698 | -97.5942
3.4735 | -97.6613
3.4771 | -97.7404
3.4808 | -97.817
3.4845 | -97.8669
3.4881 | -97.8823
3.4918 | -97.8587
3.4954 | -97.8346
3.4991 | -97.8311
3.5028 | -97.8538
3.5064 | -97.9085
3.5101 | -97.9844
3.5137 | -98.0833
3.5174 | -98.1693
3.5211 | -98.2339
3.5247 | -98.2796
3.5284 | -98.295
3.532 | -98.2659
3.5357 | -98.2157
3.5394 | -98.1543
3.543 | -98.0942
3.5467 | -98.039
3.5503 | -97.9861
3.554 | -97.9594
3.5577 | -97.9624
3.5613 | -97.974
3.565 | -97.9915
3.5687 | -98.0089
3.5723 | -98.0327
3.576 | -98.0538
3.5796 | -98.0709
3.5833 | -98.0816
3.587 | -98.0723
3.5906 | -98.0311
3.5943 | -97.9722
3.5979 | -97.9238
3.6016 | -97.9229
3.6053 | -97.979
3.6089 | -98.0955
3.6126 | -98.2534
3.6162 | -98.4216
3.6199 | -98.551
3.6236 | -98.6075
3.6272 | -98.5796
3.6309 | -98.4926
3.6345 | -98.3727
3.6382 | -98.2505
3.6419 | -98.1571
3.6455 | -98.0913
3.6492 | -98.0485
3.6528 | -98.0175
3.6565 | -97.9983
3.6602 | -97.9888
3.6638 | -97.9736
3.6675 | -97.9518
3.6711 | -97.9263
3.6748 | -97.9048
3.6785 | -97.8955
3.6821 | -97.9057
3.6858 | -97.9299
3.6894 | -97.972
3.6931 | -98.065
3.6968 | -98.152
3.7004 | -98.2575
3.7041 | -98.3796
3.7077 | -98.5069
3.7114 | -98.6333
3.7151 | -98.7348
3.7187 | -98.8033
3.7224 | -98.8418
3.726 | -98.8588
3.7297 | -98.8585
3.7334 | -98.8464
3.737 | -98.8393
3.7407 | -98.8431
3.7443 | -98.8536
3.748 | -98.8591
3.7517 | -98.8457
3.7553 | -98.8085
3.759 | -98.7545
3.7626 | -98.6747
3.7663 | -98.5749
3.77 | -98.4691
3.7736 | -98.3558
3.7773 | -98.2516
3.7809 | -98.169
3.7846 | -98.1186
3.7883 | -98.1103
3.7919 | -98.1462
3.7956 | -98.2284
3.7992 | -98.3261
3.8029 | -98.4116
3.8066 | -98.4708
3.8102 | -98.501
3.8139 | -98.4995
3.8175 | -98.4797
3.8212 | -98.456
3.8249 | -98.433
3.8285 | -98.4302
3.8322 | -98.4427
3.8358 | -98.4693
3.8395 | -98.4942
3.8432 | -98.514
3.8468 | -98.5268
3.8505 | -98.5358
3.8541 | -98.5617
3.8578 | -98.6007
3.8615 | -98.6546
3.8651 | -98.729
3.8688 | -98.8066
3.8724 | -98.8733
3.8761 | -98.92
3.8798 | -98.935
3.8834 | -98.9251
3.8871 | -98.901
3.8907 | -98.8537
3.8944 | -98.8023
3.8981 | -98.7566
3.9017 | -98.732
3.9054 | -98.7254
3.909 | -98.7365
3.9127 | -98.7647
3.9164 | -98.8057
3.92 | -98.8639
3.9237 | -98.9436
3.9273 | -99.0238
3.931 | -99.0938
3.9347 | -99.1447
3.9383 | -99.168
3.942 | -99.1567
3.9456 | -99.1123
3.9493 | -99.0488
3.953 | -98.9732
3.9566 | -98.9
3.9603 | -98.8513
3.9639 | -98.8239
3.9676 | -98.8232
3.9713 | -98.8584
3.9749 | -98.9304
3.9786 | -99.0298
3.9822 | -99.1473
3.9859 | -99.2596
3.9896 | -99.3618
3.9932 | -99.4499
3.9969 | -99.5073
4.0005 | -99.5132
4.0042 | -99.465
4.0079 | -99.377
4.0115 | -99.2678
4.0152 | -99.1537
4.0188 | -99.045
4.0225 | -98.9418
4.0262 | -98.8566
4.0298 | -98.8071
4.0335 | -98.7884
4.0372 | -98.792
4.0408 | -98.8167
4.0445 | -98.859
4.0481 | -98.9196
4.0518 | -98.988
4.0555 | -99.054
4.0591 | -99.1158
4.0628 | -99.1767
4.0664 | -99.2321
4.0701 | -99.2618
4.0738 | -99.2657
4.0774 | -99.2599
4.0811 | -99.2416
4.0847 | -99.2012
4.0884 | -99.1482
4.0921 | -99.0712
4.0957 | -98.9917
4.0994 | -98.9239
4.103 | -98.8816
4.1067 | -98.8685
4.1104 | -98.8858
4.114 | -98.9299
4.1177 | -98.9844
4.1213 | -99.0447
4.125 | -99.0934
4.1287 | -99.1126
4.1323 | -99.1015
4.136 | -99.0634
4.1396 | -99.0095
4.1433 | -98.9667
4.147 | -98.9455
4.1506 | -98.945
4.1543 | -98.9715
4.1579 | -99.0179
4.1616 | -99.0882
4.1653 | -99.1589
4.1689 | -99.242
4.1726 | -99.3343
4.1762 | -99.4259
4.1799 | -99.5227
4.1836 | -99.6035
4.1872 | -99.6582
4.1909 | -99.6852
4.1945 | -99.6811
4.1982 | -99.6307
4.2019 | -99.5386
4.2055 | -99.4177
4.2092 | -99.2853
4.2128 | -99.1512
4.2165 | -99.0268
4.2202 | -98.9174
4.2238 | -98.8781
4.2275 | -98.831
4.2311 | -98.8197
4.2348 | -98.8325
4.2385 | -98.876
4.2421 | -98.9461
4.2458 | -99.0326
4.2494 | -99.1314
4.2531 | -99.2332
4.2568 | -99.3307
4.2604 | -99.4106
4.2641 | -99.4651
4.2677 | -99.4969
4.2714 | -99.5072
4.2751 | -99.4998
4.2787 | -99.481
4.2824 | -99.4604
4.286 | -99.4497
4.2897 | -99.444
4.2934 | -99.4584
4.297 | -99.491
4.3007 | -99.5387
4.3043 | -99.6094
4.308 | -99.6908
4.3117 | -99.7761
4.3153 | -99.8594
4.319 | -99.9385
4.3226 | -100.0034
4.3263 | -100.0382
4.33 | -100.0469
4.3336 | -100.03
4.3373 | -99.9816
4.3409 | -99.911
4.3446 | -99.8279
4.3483 | -99.7476
4.3519 | -99.6698
4.3556 | -99.5864
4.3592 | -99.5105
4.3629 | -99.444
4.3666 | -99.3877
4.3702 | -99.354
4.3739 | -99.336
4.3775 | -99.3388
4.3812 | -99.3638
4.3849 | -99.394
4.3885 | -99.4267
4.3922 | -99.465
4.3958 | -99.5038
4.3995 | -99.5423
4.4032 | -99.5823
4.4068 | -99.6096
4.4105 | -99.6198
4.4141 | -99.6164
4.4178 | -99.5889
4.4215 | -99.5463
4.4251 | -99.5074
4.4288 | -99.4709
4.4324 | -99.4439
4.4361 | -99.4317
4.4398 | -99.4262
4.4434 | -99.4267
4.4471 | -99.4378
4.4507 | -99.4579
4.4544 | -99.4884
4.4581 | -99.5338
4.4617 | -99.6001
4.4654 | -99.6809
4.469 | -99.7594
4.4727 | -99.8264
4.4764 | -99.8863
4.48 | -99.923
4.4837 | -99.9265
4.4874 | -99.8966
4.491 | -99.842
4.4947 | -99.7672
4.4983 | -99.6842
4.502 | -99.6105
4.5057 | -99.5648
4.5093 | -99.5605
4.513 | -99.5914
4.5166 | -99.6511
4.5203 | -99.7338
4.524 | -99.824
4.5276 | -99.913
4.5313 | -99.9892
4.5349 | -100.0479
4.5386 | -100.0847
4.5423 | -100.1067
4.5459 | -100.1022
4.5496 | -100.0735
4.5532 | -100.028
4.5569 | -99.974
4.5606 | -99.9207
4.5642 | -99.8752
4.5679 | -99.8429
4.5715 | -99.8178
4.5752 | -99.7902
4.5789 | -99.7612
4.5825 | -99.742
4.5862 | -99.7166
4.5898 | -99.6904
4.5935 | -99.6718
4.5972 | -99.6516
4.6008 | -99.6385
4.6045 | -99.6434
4.6081 | -99.6667
4.6118 | -99.6967
4.6155 | -99.7406
4.6191 | -99.7957
4.6228 | -99.852
4.6264 | -99.9121
4.6301 | -99.97
4.6338 | -100.0288
4.6374 | -100.0849
4.6411 | -100.1366
4.6447 | -100.1764
4.6484 | -100.195
4.6521 | -100.1896
4.6557 | -100.168
4.6594 | -100.1311
4.663 | -100.0882
4.6667 | -100.0496
4.6704 | -100.0219
4.674 | -100.016
4.6777 | -100.0277
4.6813 | -100.0537
4.685 | -100.0974
4.6887 | -100.1449
4.6923 | -100.1828
4.696 | -100.197
4.6996 | -100.1804
4.7033 | -100.1343
4.707 | -100.0791
4.7106 | -100.0089
4.7143 | -99.9303
4.7179 | -99.8475
4.7216 | -99.7711
4.7253 | -99.7134
4.7289 | -99.6787
4.7326 | -99.6707
4.7362 | -99.6962
4.7399 | -99.761
4.7436 | -99.8558
4.7472 | -99.9806
4.7509 | -100.1255
4.7545 | -100.2532
4.7582 | -100.3635
4.7619 | -100.4668
4.7655 | -100.5497
4.7692 | -100.6056
4.7728 | -100.6309
4.7765 | -100.629
4.7802 | -100.5988
4.7838 | -100.5556
4.7875 | -100.508
4.7911 | -100.4623
4.7948 | -100.4246
4.7985 | -100.393
4.8021 | -100.3741
4.8058 | -100.3682
4.8094 | -100.3715
4.8131 | -100.3741
4.8168 | -100.3817
4.8204 | -100.4071
4.8241 | -100.4459
4.8277 | -100.494
4.8314 | -100.5551
4.8351 | -100.6247
4.8387 | -100.6877
4.8424 | -100.7535
4.846 | -100.8141
4.8497 | -100.8726
4.8534 | -100.9141
4.857 | -100.9363
4.8607 | -100.9486
4.8643 | -100.9517
4.868 | -100.9477
4.8717 | -100.9389
4.8753 | -100.9346
4.879 | -100.9411
4.8826 | -100.9634
4.8863 | -100.9876
4.89 | -101.0215
4.8936 | -101.0577
4.8973 | -101.0741
4.9009 | -101.078
4.9046 | -101.0809
4.9083 | -101.0705
4.9119 | -101.0577
4.9156 | -101.043
4.9192 | -101.0225
4.9229 | -100.9914
4.9266 | -100.9533
4.9302 | -100.9162
4.9339 | -100.877
4.9375 | -100.8311
4.9412 | -100.7902
4.9449 | -100.7486
4.9485 | -100.7083
4.9522 | -100.6716
4.9558 | -100.6457
4.9595 | -100.6285
4.9632 | -100.6312
4.9668 | -100.6457
4.9705 | -100.667
4.9741 | -100.7081
4.9778 | -100.7587
4.9815 | -100.8143
4.9851 | -100.8744
4.9888 | -100.9331
4.9925 | -100.9896
4.9961 | -101.0374
4.9998 | -101.0751
5.0034 | -101.0923
5.0071 | -101.0956
5.0108 | -101.0815
5.0144 | -101.0594
5.0181 | -101.0171
5.0217 | -100.9661
5.0254 | -100.9157
5.0291 | -100.8663
5.0327 | -100.8154
5.0364 | -100.7719
5.04 | -100.7366
5.0437 | -100.7121
5.0474 | -100.7032
5.051 | -100.6985
5.0547 | -100.6972
5.0583 | -100.6973
5.062 | -100.6995
5.0657 | -100.6941
5.0693 | -100.6875
5.073 | -100.6864
5.0766 | -100.6824
5.0803 | -100.6719
5.084 | -100.6566
5.0876 | -100.6517
5.0913 | -100.6524
5.0949 | -100.6614
5.0986 | -100.6817
5.1023 | -100.7035
5.1059 | -100.7248
5.1096 | -100.744
5.1132 | -100.7661
5.1169 | -100.7812
5.1206 | -100.7924
5.1242 | -100.8068
5.1279 | -100.8223
5.1315 | -100.8484
5.1352 | -100.8805
5.1389 | -100.9135
5.1425 | -100.9484
5.1462 | -100.9801
5.1498 | -100.9965
5.1535 | -101.003
5.1572 | -100.9892
5.1608 | -100.9601
5.1645 | -100.9241
5.1681 | -100.8821
5.1718 | -100.8403
5.1755 | -100.7979
5.1791 | -100.7704
5.1828 | -100.7579
5.1864 | -100.7625
5.1901 | -100.7813
5.1938 | -100.8088
5.1974 | -100.8403
5.2011 | -100.8707
5.2047 | -100.903
5.2084 | -100.9303
5.2121 | -100.9546
5.2157 | -100.9653
5.2194 | -100.9608
5.223 | -100.9379
5.2267 | -100.9076
5.2304 | -100.8758
5.234 | -100.854
5.2377 | -100.8362
5.2413 | -100.8144
5.245 | -100.809
5.2487 | -100.8155
5.2523 | -100.8214
5.256 | -100.8109
5.2596 | -100.7905
5.2633 | -100.7528
5.267 | -100.6956
5.2706 | -100.6388
5.2743 | -100.5877
5.2779 | -100.5751
5.2816 | -100.5601
5.2853 | -100.56
5.2889 | -100.5769
5.2926 | -100.6071
5.2962 | -100.6523
5.2999 | -100.7052
5.3036 | -100.7654
5.3072 | -100.8266
5.3109 | -100.885
5.3145 | -100.946
5.3182 | -101.0084
5.3219 | -101.0781
5.3255 | -101.1482
5.3292 | -101.2172
5.3328 | -101.2902
5.3365 | -101.3578
5.3402 | -101.4156
5.3438 | -101.472
5.3475 | -101.5198
5.3511 | -101.5453
5.3548 | -101.5489
5.3585 | -101.5359
5.3621 | -101.498
5.3658 | -101.4309
5.3694 | -101.3469
5.3731 | -101.2546
5.3768 | -101.1592
5.3804 | -101.0711
5.3841 | -100.9908
5.3877 | -100.925
5.3914 | -100.8789
5.3951 | -100.8483
5.3987 | -100.8306
5.4024 | -100.8322
5.406 | -100.8506
5.4097 | -100.8789
5.4134 | -100.921
5.417 | -100.9663
5.4207 | -101.0253
5.4243 | -101.0879
5.428 | -101.1468
5.4317 | -101.2114
5.4353 | -101.277
5.439 | -101.3415
5.4427 | -101.404
5.4463 | -101.4594
5.45 | -101.5184
5.4536 | -101.5587
5.4573 | -101.5721
5.461 | -101.5659
5.4646 | -101.5358
5.4683 | -101.4886
5.4719 | -101.4206
5.4756 | -101.349
5.4793 | -101.284
5.4829 | -101.2258
5.4866 | -101.1753
5.4902 | -101.1362
5.4939 | -101.1184
5.4976 | -101.1205
5.5012 | -101.1345
5.5049 | -101.1531
5.5085 | -101.1726
5.5122 | -101.1917
5.5159 | -101.1966
5.5195 | -101.186
5.5232 | -101.1701
5.5268 | -101.1469
5.5305 | -101.1287
5.5342 | -101.1043
5.5378 | -101.0869
5.5415 | -101.0797
5.5451 | -101.077
5.5488 | -101.0817
5.5525 | -101.0922
5.5561 | -101.1087
5.5598 | -101.1263
5.5634 | -101.1437
5.5671 | -101.162
5.5708 | -101.1841
5.5744 | -101.206
5.5781 | -101.2255
5.5817 | -101.2413
5.5854 | -101.2668
5.5891 | -101.2988
5.5927 | -101.3386
5.5964 | -101.3783
5.6 | -101.4267
5.6037 | -101.4793
5.6074 | -101.5202
5.611 | -101.5578
5.6147 | -101.5905
5.6183 | -101.6207
5.622 | -101.6414
5.6257 | -101.6442
5.6293 | -101.6263
5.633 | -101.593
5.6366 | -101.5413
5.6403 | -101.4788
5.644 | -101.4089
5.6476 | -101.3365
5.6513 | -101.2704
5.6549 | -101.2053
5.6586 | -101.15
5.6623 | -101.1049
5.6659 | -101.0714
5.6696 | -101.0431
5.6732 | -101.0261
5.6769 | -101.0212
5.6806 | -101.0321
5.6842 | -101.0473
5.6879 | -101.0744
5.6915 | -101.111
5.6952 | -101.1488
5.6989 | -101.187
5.7025 | -101.2189
5.7062 | -101.2436
5.7098 | -101.2678
5.7135 | -101.293
5.7172 | -101.3099
5.7208 | -101.3306
5.7245 | -101.3568
5.7281 | -101.3876
5.7318 | -101.4288
5.7355 | -101.4712
5.7391 | -101.5212
5.7428 | -101.5714
5.7464 | -101.618
5.7501 | -101.6642
5.7538 | -101.7156
5.7574 | -101.7614
5.7611 | -101.8093
5.7647 | -101.8612
5.7684 | -101.9146
5.7721 | -101.9577
5.7757 | -101.998
5.7794 | -102.0341
5.783 | -102.0589
5.7867 | -102.0759
5.7904 | -102.0763
5.794 | -102.0671
5.7977 | -102.0453
5.8013 | -102.0117
5.805 | -101.9469
5.8087 | -101.8984
5.8123 | -101.8424
5.816 | -101.7878
5.8196 | -101.7371
5.8233 | -101.693
5.827 | -101.6528
5.8306 | -101.6234
5.8343 | -101.6103
5.8379 | -101.606
5.8416 | -101.602
5.8453 | -101.6064
5.8489 | -101.6032
5.8526 | -101.5908
5.8562 | -101.576
5.8599 | -101.5538
5.8636 | -101.5247
5.8672 | -101.493
5.8709 | -101.4565
5.8745 | -101.4276
5.8782 | -101.4089
5.8819 | -101.3941
5.8855 | -101.387
5.8892 | -101.3946
5.8928 | -101.421
5.8965 | -101.4593
5.9002 | -101.5106
5.9038 | -101.5735
5.9075 | -101.644
5.9112 | -101.7171
5.9148 | -101.7932
5.9185 | -101.8647
5.9221 | -101.9278
5.9258 | -101.9818
5.9295 | -102.0218
5.9331 | -102.0523
5.9368 | -102.0642
5.9404 | -102.0729
5.9441 | -102.07
5.9478 | -102.049
5.9514 | -102.0202
5.9551 | -101.9788
5.9587 | -101.9357
5.9624 | -101.8969
5.9661 | -101.8512
5.9697 | -101.8019
5.9734 | -101.7518
5.977 | -101.7109
5.9807 | -101.6756
5.9844 | -101.6377
5.988 | -101.6094
5.9917 | -101.5867
5.9953 | -101.5657
5.999 | -101.5461
6.0027 | -101.5323
6.0063 | -101.524
6.01 | -101.5219
6.0136 | -101.527
6.0173 | -101.5419
6.021 | -101.5702
6.0246 | -101.6155
6.0283 | -101.6752
6.0319 | -101.7388
6.0356 | -101.812
6.0393 | -101.8941
6.0429 | -101.9755
6.0466 | -102.0549
6.0502 | -102.1239
6.0539 | -102.1821
6.0576 | -102.2228
6.0612 | -102.2551
6.0649 | -102.2788
6.0685 | -102.2839
6.0722 | -102.2803
6.0759 | -102.2719
6.0795 | -102.2549
6.0832 | -102.2287
6.0868 | -102.2034
6.0905 | -102.1721
6.0942 | -102.1419
6.0978 | -102.1142
6.1015 | -102.0925
6.1051 | -102.0731
6.1088 | -102.0527
6.1125 | -102.0475
6.1161 | -102.0495
6.1198 | -102.0619
6.1234 | -102.085
6.1271 | -102.1107
6.1308 | -102.1282
6.1344 | -102.1441
6.1381 | -102.1557
6.1417 | -102.1549
6.1454 | -102.1435
6.1491 | -102.1352
6.1527 | -102.1161
6.1564 | -102.0841
6.16 | -102.0474
6.1637 | -101.9989
6.1674 | -101.9492
6.171 | -101.9012
6.1747 | -101.8569
6.1783 | -101.8103
6.182 | -101.7657
6.1857 | -101.716
6.1893 | -101.6676
6.193 | -101.6254
6.1966 | -101.5891
6.2003 | -101.564
6.204 | -101.5455
6.2076 | -101.5412
6.2113 | -101.5567
6.2149 | -101.5873
6.2186 | -101.6372
6.2223 | -101.7078
6.2259 | -101.7829
6.2296 | -101.8689
6.2332 | -101.9585
6.2369 | -102.0436
6.2406 | -102.1154
6.2442 | -102.18
6.2479 | -102.2308
6.2515 | -102.2589
6.2552 | -102.2697
6.2589 | -102.2565
6.2625 | -102.2216
6.2662 | -102.177
6.2698 | -102.1264
6.2735 | -102.0759
6.2772 | -102.0349
6.2808 | -102.0036
6.2845 | -101.9868
6.2881 | -101.983
6.2918 | -101.9931
6.2955 | -102.0168
6.2991 | -102.0465
6.3028 | -102.0834
6.3064 | -102.1292
6.3101 | -102.1731
6.3138 | -102.2112
6.3174 | -102.2461
6.3211 | -102.2626
6.3247 | -102.2698
6.3284 | -102.2638
6.3321 | -102.2438
6.3357 | -102.2057
6.3394 | -102.173
6.343 | -102.1456
6.3467 | -102.1248
6.3504 | -102.1093
6.354 | -102.0994
6.3577 | -102.1008
6.3613 | -102.1074
6.365 | -102.1266
6.3687 | -102.1539
6.3723 | -102.1825
6.376 | -102.2151
6.3796 | -102.2442
6.3833 | -102.277
6.387 | -102.3076
6.3906 | -102.336
6.3943 | -102.3653
6.3979 | -102.3866
6.4016 | -102.4096
6.4053 | -102.4301
6.4089 | -102.4497
6.4126 | -102.4731
6.4163 | -102.4991
6.4199 | -102.5216
6.4236 | -102.5428
6.4272 | -102.5648
6.4309 | -102.5775
6.4346 | -102.5874
6.4382 | -102.5974
6.4419 | -102.5967
6.4455 | -102.5886
6.4492 | -102.5768
6.4529 | -102.5674
6.4565 | -102.5572
6.4602 | -102.5513
6.4638 | -102.5486
6.4675 | -102.5574
6.4712 | -102.5748
6.4748 | -102.5962
6.4785 | -102.6286
6.4821 | -102.6623
6.4858 | -102.6991
6.4895 | -102.7357
6.4931 | -102.7614
6.4968 | -102.7786
6.5004 | -102.7723
6.5041 | -102.7581
6.5078 | -102.7228
6.5114 | -102.6643
6.5151 | -102.5949
6.5187 | -102.5185
6.5224 | -102.4399
6.5261 | -102.364
6.5297 | -102.305
6.5334 | -102.2589
6.537 | -102.2265
6.5407 | -102.2198
6.5444 | -102.2359
6.548 | -102.2766
6.5517 | -102.3336
6.5553 | -102.4043
6.559 | -102.4841
6.5627 | -102.5709
6.5663 | -102.6507
6.57 | -102.7183
6.5736 | -102.7766
6.5773 | -102.8174
6.581 | -102.8451
6.5846 | -102.8532
6.5883 | -102.8449
6.5919 | -102.8264
6.5956 | -102.7988
6.5993 | -102.7664
6.6029 | -102.724
6.6066 | -102.6882
6.6102 | -102.6508
6.6139 | -102.6112
6.6176 | -102.5793
6.6212 | -102.5559
6.6249 | -102.5399
6.6285 | -102.5285
6.6322 | -102.5288
6.6359 | -102.5405
6.6395 | -102.5605
6.6432 | -102.5859
6.6468 | -102.6137
6.6505 | -102.65
6.6542 | -102.6885
6.6578 | -102.7275
6.6615 | -102.7654
6.6651 | -102.8039
6.6688 | -102.8425
6.6725 | -102.8793
6.6761 | -102.9104
6.6798 | -102.9365
6.6834 | -102.9614
6.6871 | -102.9883
6.6908 | -103.0151
6.6944 | -103.0369
6.6981 | -103.0546
6.7017 | -103.0672
6.7054 | -103.0724
6.7091 | -103.0744
6.7127 | -103.0742
6.7164 | -103.0631
6.72 | -103.0394
6.7237 | -103.0116
6.7274 | -102.9807
6.731 | -102.9422
6.7347 | -102.9047
6.7383 | -102.8668
6.742 | -102.8337
6.7457 | -102.8025
6.7493 | -102.7775
6.753 | -102.7567
6.7566 | -102.741
6.7603 | -102.7279
6.764 | -102.721
6.7676 | -102.7164
6.7713 | -102.7051
6.7749 | -102.6982
6.7786 | -102.6955
6.7823 | -102.6891
6.7859 | -102.6778
6.7896 | -102.6667
6.7932 | -102.6596
6.7969 | -102.6502
6.8006 | -102.6403
6.8042 | -102.6244
6.8079 | -102.6094
6.8115 | -102.5911
6.8152 | -102.5721
6.8189 | -102.5588
6.8225 | -102.5486
6.8262 | -102.5427
6.8298 | -102.5436
6.8335 | -102.5429
6.8372 | -102.5444
6.8408 | -102.5529
6.8445 | -102.5625
6.8481 | -102.5709
6.8518 | -102.5824
6.8555 | -102.5916
6.8591 | -102.6014
6.8628 | -102.6144
6.8665 | -102.6263
6.8701 | -102.6429
6.8738 | -102.6646
6.8774 | -102.6853
6.8811 | -102.7041
6.8848 | -102.7208
6.8884 | -102.736
6.8921 | -102.7492
6.8957 | -102.7578
6.8994 | -102.77
6.9031 | -102.7835
6.9067 | -102.8008
6.9104 | -102.8157
6.914 | -102.8352
6.9177 | -102.8594
6.9214 | -102.8829
6.925 | -102.9035
6.9287 | -102.9287
6.9323 | -102.9534
6.936 | -102.9842
6.9397 | -103.0147
6.9433 | -103.0407
6.947 | -103.0687
6.9506 | -103.0936
6.9543 | -103.1219
6.958 | -103.1523
6.9616 | -103.1814
6.9653 | -103.2082
6.9689 | -103.2266
6.9726 | -103.2484
6.9763 | -103.2674
6.9799 | -103.2835
6.9836 | -103.2961
6.9872 | -103.3107
6.9909 | -103.3311
6.9946 | -103.3496
6.9982 | -103.3631
7.0019 | -103.3734
7.0055 | -103.3802
7.0092 | -103.387
7.0129 | -103.3934
7.0165 | -103.3934
7.0202 | -103.3835
7.0238 | -103.3676
7.0275 | -103.3448
7.0312 | -103.32
7.0348 | -103.2885
7.0385 | -103.2521
7.0421 | -103.2085
7.0458 | -103.1668
7.0495 | -103.1294
7.0531 | -103.0959
7.0568 | -103.0727
7.0604 | -103.0495
7.0641 | -103.034
7.0678 | -103.0298
7.0714 | -103.0337
7.0751 | -103.048
7.0787 | -103.0717
7.0824 | -103.1054
7.0861 | -103.1442
7.0897 | -103.1933
7.0934 | -103.2438
7.097 | -103.2937
7.1007 | -103.3441
7.1044 | -103.3916
7.108 | -103.4291
7.1117 | -103.4554
7.1153 | -103.4685
7.119 | -103.4695
7.1227 | -103.4556
7.1263 | -103.4286
7.13 | -103.3896
7.1336 | -103.3393
7.1373 | -103.278
7.141 | -103.2166
7.1446 | -103.157
7.1483 | -103.1042
7.1519 | -103.0563
7.1556 | -103.0184
7.1593 | -102.9894
7.1629 | -102.973
7.1666 | -102.9743
7.1702 | -102.9859
7.1739 | -103.007
7.1776 | -103.0411
7.1812 | -103.0824
7.1849 | -103.1214
7.1885 | -103.1603
7.1922 | -103.2011
7.1959 | -103.2426
7.1995 | -103.2799
7.2032 | -103.3125
7.2068 | -103.3403
7.2105 | -103.3599
7.2142 | -103.3777
7.2178 | -103.3868
7.2215 | -103.3817
7.2251 | -103.3706
7.2288 | -103.3658
7.2325 | -103.3593
7.2361 | -103.3474
7.2398 | -103.3389
7.2434 | -103.3334
7.2471 | -103.3277
7.2508 | -103.3245
7.2544 | -103.3154
7.2581 | -103.3048
7.2617 | -103.2859
7.2654 | -103.2644
7.2691 | -103.2348
7.2727 | -103.1989
7.2764 | -103.1699
7.28 | -103.1378
7.2837 | -103.102
7.2874 | -103.0729
7.291 | -103.0509
7.2947 | -103.0338
7.2983 | -103.0214
7.302 | -103.0061
7.3057 | -102.9949
7.3093 | -102.9929
7.313 | -102.9959
7.3166 | -103.008
7.3203 | -103.0214
7.324 | -103.0436
7.3276 | -103.0776
7.3313 | -103.1145
7.335 | -103.1482
7.3386 | -103.188
7.3423 | -103.2313
7.3459 | -103.2728
7.3496 | -103.3067
7.3533 | -103.335
7.3569 | -103.3607
7.3606 | -103.3865
7.3642 | -103.4056
7.3679 | -103.4226
7.3716 | -103.4304
7.3752 | -103.4377
7.3789 | -103.4475
7.3825 | -103.4434
7.3862 | -103.4327
7.3899 | -103.4014
7.3935 | -103.3819
7.3972 | -103.3524
7.4008 | -103.321
7.4045 | -103.2858
7.4082 | -103.2443
7.4118 | -103.2034
7.4155 | -103.1607
7.4191 | -103.1131
7.4228 | -103.0685
7.4265 | -103.0321
7.4301 | -103.0051
7.4338 | -102.9807
7.4374 | -102.9649
7.4411 | -102.9628
7.4448 | -102.9736
7.4484 | -102.9916
7.4521 | -103.0222
7.4557 | -103.0651
7.4594 | -103.115
7.4631 | -103.1698
7.4667 | -103.2329
7.4704 | -103.2999
7.474 | -103.368
7.4777 | -103.4299
7.4814 | -103.4912
7.485 | -103.5392
7.4887 | -103.5764
7.4923 | -103.6051
7.496 | -103.6286
7.4997 | -103.6356
7.5033 | -103.6316
7.507 | -103.6198
7.5106 | -103.6016
7.5143 | -103.574
7.518 | -103.5526
7.5216 | -103.5351
7.5253 | -103.5214
7.5289 | -103.5154
7.5326 | -103.5165
7.5363 | -103.52
7.5399 | -103.5258
7.5436 | -103.5382
7.5472 | -103.5555
7.5509 | -103.5743
7.5546 | -103.5929
7.5582 | -103.6122
7.5619 | -103.6305
7.5655 | -103.6428
7.5692 | -103.6564
7.5729 | -103.6745
7.5765 | -103.6887
7.5802 | -103.6921
7.5838 | -103.6964
7.5875 | -103.6993
7.5912 | -103.7037
7.5948 | -103.704
7.5985 | -103.7066
7.6021 | -103.7095
7.6058 | -103.7166
7.6095 | -103.7217
7.6131 | -103.7268
7.6168 | -103.7354
7.6204 | -103.7426
7.6241 | -103.7467
7.6278 | -103.7548
7.6314 | -103.762
7.6351 | -103.7634
7.6387 | -103.7651
7.6424 | -103.7688
7.6461 | -103.7711
7.6497 | -103.7754
7.6534 | -103.7751
7.657 | -103.7777
7.6607 | -103.7833
7.6644 | -103.7876
7.668 | -103.7871
7.6717 | -103.7923
7.6753 | -103.7984
7.679 | -103.8066
7.6827 | -103.8087
7.6863 | -103.8093
7.69 | -103.8095
7.6936 | -103.8147
7.6973 | -103.8167
7.701 | -103.8173
7.7046 | -103.8154
7.7083 | -103.8058
7.7119 | -103.7918
7.7156 | -103.771
7.7193 | -103.7468
7.7229 | -103.7276
7.7266 | -103.7127
7.7302 | -103.6936
7.7339 | -103.6722
7.7376 | -103.6547
7.7412 | -103.6315
7.7449 | -103.6075
7.7485 | -103.5947
7.7522 | -103.5865
7.7559 | -103.5842
7.7595 | -103.5854
7.7632 | -103.5851
7.7668 | -103.5933
7.7705 | -103.5999
7.7742 | -103.6047
7.7778 | -103.609
7.7815 | -103.6168
7.7851 | -103.6197
7.7888 | -103.6234
7.7925 | -103.6293
7.7961 | -103.6348
7.7998 | -103.637
7.8034 | -103.6418
7.8071 | -103.6432
7.8108 | -103.6411
7.8144 | -103.6361
7.8181 | -103.6337
7.8217 | -103.6356
7.8254 | -103.6338
7.8291 | -103.6257
7.8327 | -103.6183
7.8364 | -103.602
7.8401 | -103.5813
7.8437 | -103.5661
7.8474 | -103.553
7.851 | -103.5407
7.8547 | -103.5291
7.8584 | -103.5165
7.862 | -103.4995
7.8657 | -103.4827
7.8693 | -103.466
7.873 | -103.4559
7.8767 | -103.452
7.8803 | -103.4517
7.884 | -103.4585
7.8876 | -103.4749
7.8913 | -103.4963
7.895 | -103.5206
7.8986 | -103.5468
7.9023 | -103.5798
7.9059 | -103.6173
7.9096 | -103.6548
7.9133 | -103.6945
7.9169 | -103.7321
7.9206 | -103.7693
7.9242 | -103.8049
7.9279 | -103.8439
7.9316 | -103.8869
7.9352 | -103.9251
7.9389 | -103.9643
7.9425 | -104.0112
7.9462 | -104.0607
7.9499 | -104.1039
7.9535 | -104.1483
7.9572 | -104.1906
7.9608 | -104.2273
7.9645 | -104.263
7.9682 | -104.2919
7.9718 | -104.3178
7.9755 | -104.3379
7.9791 | -104.3488
7.9828 | -104.3513
7.9865 | -104.3481
7.9901 | -104.338
7.9938 | -104.3259
7.9974 | -104.3008
8.0011 | -104.267
8.0048 | -104.2269
8.0084 | -104.1831
8.0121 | -104.1405
8.0157 | -104.0975
8.0194 | -104.0459
8.0231 | -103.9888
8.0267 | -103.9343
8.0304 | -103.8853
8.034 | -103.8381
8.0377 | -103.7946
8.0414 | -103.7583
8.045 | -103.7287
8.0487 | -103.7082
8.0523 | -103.6939
8.056 | -103.6899
8.0597 | -103.6938
8.0633 | -103.7044
8.067 | -103.7322
8.0706 | -103.7691
8.0743 | -103.8129
8.078 | -103.8617
8.0816 | -103.9149
8.0853 | -103.9695
8.0889 | -104.0227
8.0926 | -104.0757
8.0963 | -104.125
8.0999 | -104.1685
8.1036 | -104.2031
8.1072 | -104.2278
8.1109 | -104.2413
8.1146 | -104.2438
8.1182 | -104.2405
8.1219 | -104.2276
8.1255 | -104.2047
8.1292 | -104.1774
8.1329 | -104.1423
8.1365 | -104.1046
8.1402 | -104.0745
8.1438 | -104.0443
8.1475 | -104.018
8.1512 | -103.9934
8.1548 | -103.9741
8.1585 | -103.9573
8.1621 | -103.9422
8.1658 | -103.9286
8.1695 | -103.92
8.1731 | -103.9096
8.1768 | -103.9013
8.1804 | -103.8952
8.1841 | -103.8914
8.1878 | -103.8844
8.1914 | -103.8787
8.1951 | -103.8743
8.1987 | -103.8729
8.2024 | -103.8721
8.2061 | -103.8727
8.2097 | -103.8787
8.2134 | -103.8901
8.217 | -103.9067
8.2207 | -103.9285
8.2244 | -103.9486
8.228 | -103.9708
8.2317 | -103.9935
8.2353 | -104.0175
8.239 | -104.0334
8.2427 | -104.0502
8.2463 | -104.0665
8.25 | -104.0741
8.2536 | -104.0769
8.2573 | -104.0778
8.261 | -104.0707
8.2646 | -104.059
8.2683 | -104.0461
8.2719 | -104.0326
8.2756 | -104.0153
8.2793 | -103.9962
8.2829 | -103.979
8.2866 | -103.9627
8.2903 | -103.9447
8.2939 | -103.9358
8.2976 | -103.9289
8.3012 | -103.9199
8.3049 | -103.9103
8.3086 | -103.9034
8.3122 | -103.9024
8.3159 | -103.9057
8.3195 | -103.9108
8.3232 | -103.9162
8.3269 | -103.9279
8.3305 | -103.948
8.3342 | -103.9686
8.3378 | -103.9927
8.3415 | -104.0207
8.3452 | -104.0564
8.3488 | -104.098
8.3525 | -104.1429
8.3561 | -104.1905
8.3598 | -104.2496
8.3635 | -104.3088
8.3671 | -104.361
8.3708 | -104.4105
8.3744 | -104.4544
8.3781 | -104.4847
8.3818 | -104.5112
8.3854 | -104.5331
8.3891 | -104.5377
8.3927 | -104.5331
8.3964 | -104.5198
8.4001 | -104.4982
8.4037 | -104.471
8.4074 | -104.4406
8.411 | -104.4021
8.4147 | -104.3615
8.4184 | -104.3243
8.422 | -104.2939
8.4257 | -104.2755
8.4293 | -104.2658
8.433 | -104.2615
8.4367 | -104.2639
8.4403 | -104.2768
8.444 | -104.3015
8.4476 | -104.3365
8.4513 | -104.3603
8.455 | -104.3804
8.4586 | -104.4003
8.4623 | -104.4245
8.4659 | -104.4458
8.4696 | -104.4595
8.4733 | -104.4649
8.4769 | -104.4633
8.4806 | -104.4517
8.4842 | -104.4305
8.4879 | -104.4068
8.4916 | -104.3822
8.4952 | -104.3495
8.4989 | -104.3148
8.5025 | -104.2826
8.5062 | -104.2439
8.5099 | -104.2071
8.5135 | -104.1751
8.5172 | -104.1417
8.5208 | -104.1137
8.5245 | -104.0922
8.5282 | -104.0726
8.5318 | -104.0535
8.5355 | -104.0454
8.5391 | -104.0399
8.5428 | -104.0337
8.5465 | -104.0319
8.5501 | -104.0256
8.5538 | -104.0194
8.5574 | -104.0177
8.5611 | -104.0174
8.5648 | -104.0144
8.5684 | -104.0033
8.5721 | -103.9943
8.5757 | -103.9846
8.5794 | -103.9669
8.5831 | -103.9476
8.5867 | -103.9246
8.5904 | -103.8948
8.594 | -103.8695
8.5977 | -103.8455
8.6014 | -103.8193
8.605 | -103.7986
8.6087 | -103.7834
8.6123 | -103.7743
8.616 | -103.7682
8.6197 | -103.7696
8.6233 | -103.7854
8.627 | -103.8063
8.6306 | -103.8354
8.6343 | -103.8746
8.638 | -103.9241
8.6416 | -103.9815
8.6453 | -104.0432
8.6489 | -104.1133
8.6526 | -104.1894
8.6563 | -104.2635
8.6599 | -104.34
8.6636 | -104.4101
8.6672 | -104.4765
8.6709 | -104.5339
8.6746 | -104.5797
8.6782 | -104.6129
8.6819 | -104.6374
8.6855 | -104.6497
8.6892 | -104.6491
8.6929 | -104.6405
8.6965 | -104.6239
8.7002 | -104.5985
8.7038 | -104.5672
8.7075 | -104.538
8.7112 | -104.5037
8.7148 | -104.4662
8.7185 | -104.4291
8.7221 | -104.3893
8.7258 | -104.3509
8.7295 | -104.3169
8.7331 | -104.2787
8.7368 | -104.2418
8.7404 | -104.2065
8.7441 | -104.1705
8.7478 | -104.1383
8.7514 | -104.0988
8.7551 | -104.0611
8.7587 | -104.0238
8.7624 | -103.987
8.7661 | -103.9508
8.7697 | -103.9164
8.7734 | -103.8879
8.7771 | -103.8576
8.7807 | -103.8316
8.7844 | -103.8155
8.788 | -103.8012
8.7917 | -103.794
8.7954 | -103.7883
8.799 | -103.7907
8.8027 | -103.7967
8.8063 | -103.8091
8.81 | -103.8201
8.8137 | -103.8341
8.8173 | -103.8542
8.821 | -103.8793
8.8246 | -103.9061
8.8283 | -103.9323
8.832 | -103.9585
8.8356 | -103.9876
8.8393 | -104.0243
8.8429 | -104.063
8.8466 | -104.0994
8.8503 | -104.1361
8.8539 | -104.1752
8.8576 | -104.2151
8.8612 | -104.2575
8.8649 | -104.3035
8.8686 | -104.3463
8.8722 | -104.3887
8.8759 | -104.4291
8.8795 | -104.473
8.8832 | -104.5111
8.8869 | -104.5446
8.8905 | -104.5784
8.8942 | -104.6136
8.8978 | -104.6498
8.9015 | -104.683
8.9052 | -104.7088
8.9088 | -104.7282
8.9125 | -104.7409
8.9161 | -104.7545
8.9198 | -104.7637
8.9235 | -104.7746
8.9271 | -104.7802
8.9308 | -104.7844
8.9344 | -104.7844
8.9381 | -104.7841
8.9418 | -104.7817
8.9454 | -104.7749
8.9491 | -104.7658
8.9527 | -104.7607
8.9564 | -104.7535
8.9601 | -104.7411
8.9637 | -104.7236
8.9674 | -104.7013
8.971 | -104.6783
8.9747 | -104.6499
8.9784 | -104.6276
8.982 | -104.6059
8.9857 | -104.5801
8.9893 | -104.5551
8.993 | -104.5337
8.9967 | -104.5086
9.0003 | -104.4903
9.004 | -104.4772
9.0076 | -104.464
9.0113 | -104.4538
9.015 | -104.4486
9.0186 | -104.4505
9.0223 | -104.4553
9.0259 | -104.4591
9.0296 | -104.4695
9.0333 | -104.4859
9.0369 | -104.5118
9.0406 | -104.5402
9.0442 | -104.5704
9.0479 | -104.6022
9.0516 | -104.6319
9.0552 | -104.6584
9.0589 | -104.6857
9.0625 | -104.7122
9.0662 | -104.7389
9.0699 | -104.765
9.0735 | -104.7911
9.0772 | -104.8126
9.0808 | -104.8264
9.0845 | -104.8327
9.0882 | -104.8337
9.0918 | -104.8315
9.0955 | -104.8265
9.0991 | -104.8192
9.1028 | -104.81
9.1065 | -104.7969
9.1101 | -104.7804
9.1138 | -104.7639
9.1174 | -104.7443
9.1211 | -104.7274
9.1248 | -104.7106
9.1284 | -104.694
9.1321 | -104.6845
9.1357 | -104.6759
9.1394 | -104.6722
9.1431 | -104.678
9.1467 | -104.6887
9.1504 | -104.7059
9.154 | -104.7248
9.1577 | -104.7429
9.1614 | -104.768
9.165 | -104.7961
9.1687 | -104.8241
9.1723 | -104.8495
9.176 | -104.873
9.1797 | -104.8957
9.1833 | -104.9127
9.187 | -104.929
9.1906 | -104.9394
9.1943 | -104.9417
9.198 | -104.9386
9.2016 | -104.932
9.2053 | -104.9211
9.2089 | -104.9061
9.2126 | -104.8962
9.2163 | -104.8876
9.2199 | -104.8794
9.2236 | -104.865
9.2272 | -104.8539
9.2309 | -104.8395
9.2346 | -104.827
9.2382 | -104.8175
9.2419 | -104.8103
9.2456 | -104.8031
9.2492 | -104.7978
9.2529 | -104.7913
9.2565 | -104.7861
9.2602 | -104.7854
9.2639 | -104.7868
9.2675 | -104.7863
9.2712 | -104.7845
9.2748 | -104.785
9.2785 | -104.7855
9.2822 | -104.788
9.2858 | -104.792
9.2895 | -104.7953
9.2931 | -104.8078
9.2968 | -104.8192
9.3005 | -104.8291
9.3041 | -104.8408
9.3078 | -104.857
9.3114 | -104.8734
9.3151 | -104.8888
9.3188 | -104.908
9.3224 | -104.925
9.3261 | -104.9404
9.3297 | -104.9582
9.3334 | -104.9752
9.3371 | -104.985
9.3407 | -104.9886
9.3444 | -104.9918
9.348 | -104.9909
9.3517 | -104.9895
9.3554 | -104.9822
9.359 | -104.9743
9.3627 | -104.9607
9.3663 | -104.9451
9.37 | -104.9263
9.3737 | -104.9065
9.3773 | -104.8854
9.381 | -104.8653
9.3846 | -104.8462
9.3883 | -104.8266
9.392 | -104.8068
9.3956 | -104.7869
9.3993 | -104.7657
9.4029 | -104.743
9.4066 | -104.7254
9.4103 | -104.7046
9.4139 | -104.685
9.4176 | -104.6683
9.4212 | -104.6502
9.4249 | -104.6295
9.4286 | -104.6049
9.4322 | -104.5887
9.4359 | -104.577
9.4395 | -104.5701
9.4432 | -104.5629
9.4469 | -104.5546
9.4505 | -104.5561
9.4542 | -104.5581
9.4578 | -104.5646
9.4615 | -104.5739
9.4652 | -104.5808
9.4688 | -104.5901
9.4725 | -104.6074
9.4761 | -104.6281
9.4798 | -104.6468
9.4835 | -104.6649
9.4871 | -104.6805
9.4908 | -104.7008
9.4944 | -104.7217
9.4981 | -104.7357
9.5018 | -104.7452
9.5054 | -104.7579
9.5091 | -104.7715
9.5127 | -104.7816
9.5164 | -104.789
9.5201 | -104.7937
9.5237 | -104.7951
9.5274 | -104.7978
9.531 | -104.7968
9.5347 | -104.7884
9.5384 | -104.7789
9.542 | -104.7699
9.5457 | -104.7547
9.5493 | -104.7371
9.553 | -104.7176
9.5567 | -104.6962
9.5603 | -104.6745
9.564 | -104.6553
9.5676 | -104.6354
9.5713 | -104.6126
9.575 | -104.5932
9.5786 | -104.5758
9.5823 | -104.5637
9.5859 | -104.5603
9.5896 | -104.5546
9.5933 | -104.5542
9.5969 | -104.558
9.6006 | -104.5679
9.6042 | -104.5851
9.6079 | -104.5995
9.6116 | -104.6169
9.6152 | -104.6353
9.6189 | -104.6521
9.6225 | -104.6742
9.6262 | -104.7011
9.6299 | -104.7268
9.6335 | -104.7509
9.6372 | -104.775
9.6408 | -104.794
9.6445 | -104.8089
9.6482 | -104.8237
9.6518 | -104.8353
9.6555 | -104.8472
9.6591 | -104.8598
9.6628 | -104.8701
9.6665 | -104.877
9.6701 | -104.8817
9.6738 | -104.887
9.6774 | -104.8938
9.6811 | -104.8988
9.6848 | -104.9055
9.6884 | -104.9152
9.6921 | -104.9254
9.6958 | -104.9363
9.6994 | -104.9492
9.7031 | -104.9638
9.7067 | -104.9795
9.7104 | -104.996
9.7141 | -105.0099
9.7177 | -105.0232
9.7214 | -105.039
9.725 | -105.0549
9.7287 | -105.0714
9.7324 | -105.0893
9.736 | -105.105
9.7397 | -105.1189
9.7433 | -105.1299
9.747 | -105.1325
9.7507 | -105.1334
9.7543 | -105.1291
9.758 | -105.1262
9.7616 | -105.1176
9.7653 | -105.1
9.769 | -105.0774
9.7726 | -105.0516
9.7763 | -105.0232
9.7799 | -104.995
9.7836 | -104.9673
9.7873 | -104.9384
9.7909 | -104.9054
9.7946 | -104.8733
9.7982 | -104.8457
9.8019 | -104.8177
9.8056 | -104.7927
9.8092 | -104.7733
9.8129 | -104.7565
9.8165 | -104.7442
9.8202 | -104.7366
9.8239 | -104.735
9.8275 | -104.7327
9.8312 | -104.7318
9.8348 | -104.7321
9.8385 | -104.7371
9.8422 | -104.7435
9.8458 | -104.7529
9.8495 | -104.7611
9.8531 | -104.768
9.8568 | -104.7752
9.8605 | -104.7827
9.8641 | -104.7918
9.8678 | -104.801
9.8714 | -104.8043
9.8751 | -104.8052
9.8788 | -104.8073
9.8824 | -104.8041
9.8861 | -104.7974
9.8897 | -104.7919
9.8934 | -104.7859
9.8971 | -104.7784
9.9007 | -104.7683
9.9044 | -104.7632
9.908 | -104.7549
9.9117 | -104.7481
9.9154 | -104.7412
9.919 | -104.7367
9.9227 | -104.7339
9.9263 | -104.7294
9.93 | -104.7283
9.9337 | -104.7297
9.9373 | -104.7307
9.941 | -104.7293
9.9446 | -104.7294
9.9483 | -104.7303
9.952 | -104.731
9.9556 | -104.7304
9.9593 | -104.7282
9.9629 | -104.7273
9.9666 | -104.7267
9.9703 | -104.7264
9.9739 | -104.7263
9.9776 | -104.7286
9.9812 | -104.7289
9.9849 | -104.7332
9.9886 | -104.7379
9.9922 | -104.7464
9.9959 | -104.7574
9.9995 | -104.7666
10.0032 | -104.7759
10.0069 | -104.7878
10.0105 | -104.8032
10.0142 | -104.8255
10.0178 | -104.8469
10.0215 | -104.8674
10.0252 | -104.8907
10.0288 | -104.912
10.0325 | -104.9247
10.0361 | -104.9449
10.0398 | -104.9614
10.0435 | -104.9739
10.0471 | -104.9835
10.0508 | -104.9883
10.0544 | -104.9962
10.0581 | -105.0007
10.0618 | -105.0034
10.0654 | -105.0052
10.0691 | -105.0002
10.0727 | -104.9969
10.0764 | -104.9941
10.0801 | -104.9854
10.0837 | -104.9776
10.0874 | -104.9739
10.091 | -104.9734
10.0947 | -104.9707
10.0984 | -104.9642
10.102 | -104.959
10.1057 | -104.9514
10.1093 | -104.9452
10.113 | -104.9427
10.1167 | -104.9417
10.1203 | -104.9409
10.124 | -104.9384
10.1276 | -104.932
10.1313 | -104.9206
10.135 | -104.9106
10.1386 | -104.8994
10.1423 | -104.8861
10.1459 | -104.8745
10.1496 | -104.8602
10.1533 | -104.8424
10.1569 | -104.826
10.1606 | -104.8084
10.1642 | -104.7907
10.1679 | -104.7723
10.1716 | -104.754
10.1752 | -104.7382
10.1789 | -104.7267
10.1826 | -104.718
10.1862 | -104.7114
10.1899 | -104.7094
10.1935 | -104.7149
10.1972 | -104.7237
10.2009 | -104.7399
10.2045 | -104.7614
10.2082 | -104.7896
10.2118 | -104.8221
10.2155 | -104.8591
10.2192 | -104.9042
10.2228 | -104.9498
10.2265 | -104.9959
10.2301 | -105.0483
10.2338 | -105.1002
10.2375 | -105.1497
10.2411 | -105.1991
10.2448 | -105.2493
10.2484 | -105.2963
10.2521 | -105.3366
10.2558 | -105.3786
10.2594 | -105.4162
10.2631 | -105.4407
10.2667 | -105.4587
10.2704 | -105.4746
10.2741 | -105.484
10.2777 | -105.4871
10.2814 | -105.4839
10.285 | -105.4691
10.2887 | -105.4492
10.2924 | -105.4234
10.296 | -105.3903
10.2997 | -105.3525
10.3033 | -105.3138
10.307 | -105.2707
10.3107 | -105.2266
10.3143 | -105.187
10.318 | -105.1504
10.3216 | -105.1141
10.3253 | -105.0819
10.329 | -105.0525
10.3326 | -105.026
10.3363 | -105.0022
10.3399 | -104.9897
10.3436 | -104.9786
10.3473 | -104.9753
10.3509 | -104.9759
10.3546 | -104.9801
10.3582 | -104.9848
10.3619 | -104.9962
10.3656 | -105.0095
10.3692 | -105.0216
10.3729 | -105.0377
10.3765 | -105.0538
10.3802 | -105.0684
10.3839 | -105.0811
10.3875 | -105.0934
10.3912 | -105.1024
10.3948 | -105.1091
10.3985 | -105.1186
10.4022 | -105.1255
10.4058 | -105.127
10.4095 | -105.127
10.4131 | -105.1302
10.4168 | -105.1352
10.4205 | -105.1353
10.4241 | -105.1324
10.4278 | -105.1313
10.4314 | -105.1317
10.4351 | -105.136
10.4388 | -105.1319
10.4424 | -105.1279
10.4461 | -105.1262
10.4497 | -105.1226
10.4534 | -105.1172
10.4571 | -105.1132
10.4607 | -105.1129
10.4644 | -105.11
10.468 | -105.1047
10.4717 | -105.0987
10.4754 | -105.0924
10.479 | -105.0875
10.4827 | -105.0821
10.4863 | -105.0794
10.49 | -105.0735
10.4937 | -105.0695
10.4973 | -105.0658
10.501 | -105.0615
10.5046 | -105.0559
10.5083 | -105.0536
10.512 | -105.0484
10.5156 | -105.044
10.5193 | -105.0425
10.5229 | -105.0413
10.5266 | -105.0416
10.5303 | -105.0471
10.5339 | -105.0553
10.5376 | -105.0644
10.5412 | -105.0767
10.5449 | -105.0914
10.5486 | -105.1081
10.5522 | -105.1293
10.5559 | -105.1518
10.5595 | -105.1804
10.5632 | -105.2027
10.5669 | -105.2254
10.5705 | -105.2462
10.5742 | -105.2702
10.5778 | -105.3
10.5815 | -105.3268
10.5852 | -105.3504
10.5888 | -105.3728
10.5925 | -105.4011
10.5961 | -105.4276
10.5998 | -105.4485
10.6035 | -105.4721
10.6071 | -105.4943
10.6108 | -105.5154
10.6144 | -105.5379
10.6181 | -105.5614
10.6218 | -105.5885
10.6254 | -105.618
10.6291 | -105.6445
10.6327 | -105.6675
10.6364 | -105.6902
10.6401 | -105.7133
10.6437 | -105.7343
10.6474 | -105.7539
10.651 | -105.7704
10.6547 | -105.784
10.6584 | -105.7977
10.662 | -105.8061
10.6657 | -105.8156
10.6694 | -105.8193
10.673 | -105.8161
10.6767 | -105.8116
10.6803 | -105.7988
10.684 | -105.7777
10.6877 | -105.7541
10.6913 | -105.7291
10.695 | -105.7024
10.6986 | -105.6689
10.7023 | -105.6368
10.706 | -105.6022
10.7096 | -105.5688
10.7133 | -105.5338
10.7169 | -105.5009
10.7206 | -105.4654
10.7243 | -105.4295
10.7279 | -105.3966
10.7316 | -105.3623
10.7352 | -105.3293
10.7389 | -105.2981
10.7426 | -105.2706
10.7462 | -105.2423
10.7499 | -105.217
10.7535 | -105.1949
10.7572 | -105.1762
10.7609 | -105.1545
10.7645 | -105.1373
10.7682 | -105.1226
10.7718 | -105.1078
10.7755 | -105.098
10.7792 | -105.0859
10.7828 | -105.0756
10.7865 | -105.0691
10.7901 | -105.0668
10.7938 | -105.0659
10.7975 | -105.0676
10.8011 | -105.0702
10.8048 | -105.0737
10.8084 | -105.0795
10.8121 | -105.0879
10.8158 | -105.0942
10.8194 | -105.1049
10.8231 | -105.1177
10.8267 | -105.1307
10.8304 | -105.1438
10.8341 | -105.1575
10.8377 | -105.176
10.8414 | -105.1931
10.845 | -105.2097
10.8487 | -105.2271
10.8524 | -105.2451
10.856 | -105.268
10.8597 | -105.2848
10.8633 | -105.3006
10.867 | -105.3134
10.8707 | -105.3276
10.8743 | -105.3434
10.878 | -105.3584
10.8816 | -105.3744
10.8853 | -105.3865
10.889 | -105.3987
10.8926 | -105.4076
10.8963 | -105.4171
10.8999 | -105.4255
10.9036 | -105.4363
10.9073 | -105.4461
10.9109 | -105.4573
10.9146 | -105.4703
10.9182 | -105.4855
10.9219 | -105.4988
10.9256 | -105.5135
10.9292 | -105.5294
10.9329 | -105.5456
10.9365 | -105.5635
10.9402 | -105.579
10.9439 | -105.5914
10.9475 | -105.601
10.9512 | -105.6119
10.9548 | -105.6193
10.9585 | -105.6202
10.9622 | -105.6232
10.9658 | -105.6258
10.9695 | -105.6236
10.9731 | -105.6211
10.9768 | -105.6127
10.9805 | -105.5993
10.9841 | -105.5843
10.9878 | -105.5678
10.9914 | -105.5486
10.9951 | -105.5285
10.9988 | -105.5082
11.0024 | -105.4852
11.0061 | -105.4595
11.0097 | -105.434
11.0134 | -105.4065
11.0171 | -105.3803
11.0207 | -105.3571
11.0244 | -105.3341
11.028 | -105.3139
11.0317 | -105.2937
11.0354 | -105.2774
11.039 | -105.263
11.0427 | -105.2527
11.0463 | -105.2454
11.05 | -105.2374
11.0537 | -105.232
11.0573 | -105.2271
11.061 | -105.2211
11.0646 | -105.2211
11.0683 | -105.2228
11.072 | -105.2219
11.0756 | -105.218
11.0793 | -105.2114
11.0829 | -105.2045
11.0866 | -105.1931
11.0903 | -105.1811
11.0939 | -105.1676
11.0976 | -105.1572
11.1012 | -105.148
11.1049 | -105.1399
11.1086 | -105.1308
11.1122 | -105.1193
11.1159 | -105.1087
11.1196 | -105.1056
11.1232 | -105.1035
11.1269 | -105.0999
11.1305 | -105.0977
11.1342 | -105.0986
11.1379 | -105.1057
11.1415 | -105.1145
11.1452 | -105.127
11.1488 | -105.1405
11.1525 | -105.1516
11.1562 | -105.1671
11.1598 | -105.1871
11.1635 | -105.2065
11.1671 | -105.2244
11.1708 | -105.2444
11.1745 | -105.2674
11.1781 | -105.2902
11.1818 | -105.311
11.1854 | -105.3344
11.1891 | -105.3558
11.1928 | -105.3778
11.1964 | -105.3974
11.2001 | -105.4168
11.2037 | -105.4382
11.2074 | -105.4581
11.2111 | -105.478
11.2147 | -105.4989
11.2184 | -105.5191
11.222 | -105.5369
11.2257 | -105.5491
11.2294 | -105.5595
11.233 | -105.5735
11.2367 | -105.5923
11.2403 | -105.6084
11.244 | -105.6212
11.2477 | -105.6325
11.2513 | -105.6437
11.255 | -105.654
11.2586 | -105.6597
11.2623 | -105.6605
11.266 | -105.6547
11.2696 | -105.6441
11.2733 | -105.6359
11.2769 | -105.6274
11.2806 | -105.6119
11.2843 | -105.5935
11.2879 | -105.5731
11.2916 | -105.5559
11.2952 | -105.5404
11.2989 | -105.5201
11.3026 | -105.4974
11.3062 | -105.4761
11.3099 | -105.4543
11.3135 | -105.4331
11.3172 | -105.4166
11.3209 | -105.3985
11.3245 | -105.3791
11.3282 | -105.3643
11.3318 | -105.3565
11.3355 | -105.3484
11.3392 | -105.3445
11.3428 | -105.3441
11.3465 | -105.3518
11.3501 | -105.3599
11.3538 | -105.3707
11.3575 | -105.3891
11.3611 | -105.4129
11.3648 | -105.4358
11.3684 | -105.4583
11.3721 | -105.484
11.3758 | -105.514
11.3794 | -105.5446
11.3831 | -105.5773
11.3867 | -105.6128
11.3904 | -105.6436
11.3941 | -105.6717
11.3977 | -105.6959
11.4014 | -105.7157
11.405 | -105.7323
11.4087 | -105.7438
11.4124 | -105.7508
11.416 | -105.7538
11.4197 | -105.7564
11.4233 | -105.7571
11.427 | -105.7547
11.4307 | -105.7523
11.4343 | -105.7496
11.438 | -105.7422
11.4416 | -105.7329
11.4453 | -105.7279
11.449 | -105.7218
11.4526 | -105.7154
11.4563 | -105.7099
11.4599 | -105.7059
11.4636 | -105.7066
11.4673 | -105.7102
11.4709 | -105.7163
11.4746 | -105.7233
11.4782 | -105.7337
11.4819 | -105.743
11.4856 | -105.753
11.4892 | -105.7671
11.4929 | -105.7827
11.4965 | -105.7954
11.5002 | -105.8096
11.5039 | -105.8225
11.5075 | -105.8373
11.5112 | -105.8481
11.5148 | -105.8609
11.5185 | -105.8738
11.5222 | -105.8849
11.5258 | -105.8972
11.5295 | -105.9109
11.5331 | -105.9268
11.5368 | -105.9416
11.5405 | -105.95
11.5441 | -105.9534
11.5478 | -105.9532
11.5514 | -105.9539
11.5551 | -105.9552
11.5588 | -105.9509
11.5624 | -105.9459
11.5661 | -105.9389
11.5697 | -105.9316
11.5734 | -105.9237
11.5771 | -105.913
11.5807 | -105.9012
11.5844 | -105.8877
11.5881 | -105.8745
11.5917 | -105.8619
11.5954 | -105.8491
11.599 | -105.8352
11.6027 | -105.824
11.6064 | -105.8152
11.61 | -105.804
11.6137 | -105.8051
11.6173 | -105.8017
11.621 | -105.8008
11.6247 | -105.8036
11.6283 | -105.8074
11.632 | -105.8099
11.6356 | -105.8137
11.6393 | -105.8191
11.643 | -105.8248
11.6466 | -105.831
11.6503 | -105.8387
11.6539 | -105.8454
11.6576 | -105.851
11.6613 | -105.8551
11.6649 | -105.8619
11.6686 | -105.8665
11.6722 | -105.8695
11.6759 | -105.8711
11.6796 | -105.8737
11.6832 | -105.8765
11.6869 | -105.8794
11.6905 | -105.8787
11.6942 | -105.878
11.6979 | -105.8778
11.7015 | -105.8773
11.7052 | -105.8763
11.7088 | -105.8808
11.7125 | -105.8877
11.7162 | -105.8932
11.7198 | -105.8977
11.7235 | -105.9053
11.7271 | -105.9188
11.7308 | -105.928
11.7345 | -105.9355
11.7381 | -105.9416
11.7418 | -105.9478
11.7454 | -105.9512
11.7491 | -105.9541
11.7528 | -105.9592
11.7564 | -105.9613
11.7601 | -105.9622
11.7637 | -105.963
11.7674 | -105.9624
11.7711 | -105.9601
11.7747 | -105.9585
11.7784 | -105.9609
11.782 | -105.961
11.7857 | -105.9582
11.7894 | -105.9558
11.793 | -105.9526
11.7967 | -105.9505
11.8003 | -105.9499
11.804 | -105.9527
11.8077 | -105.9526
11.8113 | -105.9507
11.815 | -105.9521
11.8186 | -105.9526
11.8223 | -105.951
11.826 | -105.9527
11.8296 | -105.9548
11.8333 | -105.9572
11.8369 | -105.957
11.8406 | -105.9563
11.8443 | -105.9583
11.8479 | -105.9629
11.8516 | -105.9673
11.8552 | -105.9683
11.8589 | -105.9706
11.8626 | -105.9778
11.8662 | -105.9853
11.8699 | -105.992
11.8735 | -105.998
11.8772 | -106.0088
11.8809 | -106.0206
11.8845 | -106.0287
11.8882 | -106.0404
11.8918 | -106.0497
11.8955 | -106.0554
11.8992 | -106.0629
11.9028 | -106.0692
11.9065 | -106.0764
11.9101 | -106.0814
11.9138 | -106.0822
11.9175 | -106.0876
11.9211 | -106.0919
11.9248 | -106.0914
11.9284 | -106.0888
11.9321 | -106.0814
11.9358 | -106.0723
11.9394 | -106.0593
11.9431 | -106.0404
11.9467 | -106.0208
11.9504 | -105.998
11.9541 | -105.9727
11.9577 | -105.9466
11.9614 | -105.9196
11.965 | -105.8934
11.9687 | -105.8678
11.9724 | -105.8417
11.976 | -105.8158
11.9797 | -105.7956
11.9833 | -105.7789
11.987 | -105.7618
11.9907 | -105.7453
11.9943 | -105.7335
11.998 | -105.7206
12.0016 | -105.7124
12.0053 | -105.7082
12.009 | -105.7099
12.0126 | -105.7163
12.0163 | -105.7212
12.0199 | -105.7258
12.0236 | -105.7349
12.0273 | -105.7505
12.0309 | -105.7665
12.0346 | -105.7813
12.0382 | -105.7977
12.0419 | -105.8136
12.0456 | -105.8302
12.0492 | -105.8419
12.0529 | -105.8524
12.0565 | -105.8672
12.0602 | -105.8816
12.0639 | -105.8945
12.0675 | -105.904
12.0712 | -105.9083
12.0748 | -105.9163
12.0785 | -105.9215
12.0822 | -105.9289
12.0858 | -105.9364
12.0895 | -105.9374
12.0932 | -105.935
12.0968 | -105.9341
12.1005 | -105.9349
12.1041 | -105.9391
12.1078 | -105.9413
12.1115 | -105.9438
12.1151 | -105.9472
12.1188 | -105.9471
12.1224 | -105.9454
12.1261 | -105.9441
12.1298 | -105.9409
12.1334 | -105.9397
12.1371 | -105.9388
12.1407 | -105.9353
12.1444 | -105.9294
12.1481 | -105.9252
12.1517 | -105.9196
12.1554 | -105.9103
12.159 | -105.9002
12.1627 | -105.8892
12.1664 | -105.881
12.17 | -105.8697
12.1737 | -105.8588
12.1773 | -105.8469
12.181 | -105.8329
12.1847 | -105.8161
12.1883 | -105.8028
12.192 | -105.7919
12.1956 | -105.7803
12.1993 | -105.7691
12.203 | -105.757
12.2066 | -105.7477
12.2103 | -105.7394
12.2139 | -105.7283
12.2176 | -105.72
12.2213 | -105.7155
12.2249 | -105.7155
12.2286 | -105.7165
12.2322 | -105.7165
12.2359 | -105.7138
12.2396 | -105.7163
12.2432 | -105.7233
12.2469 | -105.7294
12.2505 | -105.7334
12.2542 | -105.7376
12.2579 | -105.7411
12.2615 | -105.7446
12.2652 | -105.7499
12.2688 | -105.754
12.2725 | -105.7565
12.2762 | -105.7599
12.2798 | -105.7623
12.2835 | -105.7646
12.2871 | -105.7641
12.2908 | -105.7634
12.2945 | -105.7642
12.2981 | -105.7661
12.3018 | -105.7676
12.3054 | -105.7709
12.3091 | -105.7712
12.3128 | -105.7704
12.3164 | -105.7734
12.3201 | -105.7771
12.3237 | -105.7812
12.3274 | -105.7894
12.3311 | -105.7993
12.3347 | -105.8078
12.3384 | -105.8154
12.342 | -105.8258
12.3457 | -105.8358
12.3494 | -105.8501
12.353 | -105.8666
12.3567 | -105.8799
12.3603 | -105.8953
12.364 | -105.9109
12.3677 | -105.9252
12.3713 | -105.9396
12.375 | -105.9521
12.3786 | -105.963
12.3823 | -105.9743
12.386 | -105.984
12.3896 | -105.9928
12.3933 | -106.0046
12.3969 | -106.0102
12.4006 | -106.0082
12.4043 | -106.004
12.4079 | -105.9991
12.4116 | -105.9955
12.4152 | -105.9932
12.4189 | -105.9878
12.4226 | -105.9812
12.4262 | -105.9749
12.4299 | -105.9672
12.4335 | -105.9575
12.4372 | -105.9461
12.4409 | -105.9363
12.4445 | -105.9304
12.4482 | -105.9278
12.4518 | -105.9259
12.4555 | -105.9252
12.4592 | -105.9252
12.4628 | -105.9285
12.4665 | -105.9368
12.4701 | -105.9461
12.4738 | -105.9549
12.4775 | -105.9668
12.4811 | -105.9811
12.4848 | -105.9985
12.4884 | -106.0184
12.4921 | -106.0418
12.4958 | -106.0661
12.4994 | -106.093
12.5031 | -106.1191
12.5067 | -106.1461
12.5104 | -106.1737
12.5141 | -106.2039
12.5177 | -106.2313
12.5214 | -106.2593
12.525 | -106.2827
12.5287 | -106.3035
12.5324 | -106.3216
12.536 | -106.3402
12.5397 | -106.3577
12.5434 | -106.3747
12.547 | -106.3901
12.5507 | -106.4001
12.5543 | -106.4096
12.558 | -106.4193
12.5617 | -106.425
12.5653 | -106.429
12.569 | -106.4316
12.5726 | -106.4304
12.5763 | -106.4276
12.58 | -106.4203
12.5836 | -106.4125
12.5873 | -106.4063
12.5909 | -106.3963
12.5946 | -106.3852
12.5983 | -106.3746
12.6019 | -106.3651
12.6056 | -106.3575
12.6092 | -106.3497
12.6129 | -106.3412
12.6166 | -106.335
12.6202 | -106.3295
12.6239 | -106.3211
12.6275 | -106.3133
12.6312 | -106.3028
12.6349 | -106.2916
12.6385 | -106.2826
12.6422 | -106.2732
12.6458 | -106.2675
12.6495 | -106.2607
12.6532 | -106.2538
12.6568 | -106.2443
12.6605 | -106.2351
12.6641 | -106.226
12.6678 | -106.2187
12.6715 | -106.2102
12.6751 | -106.2008
12.6788 | -106.1896
12.6824 | -106.1807
12.6861 | -106.1754
12.6898 | -106.1678
12.6934 | -106.1624
12.6971 | -106.1589
12.7007 | -106.159
12.7044 | -106.1531
12.7081 | -106.1508
12.7117 | -106.1471
12.7154 | -106.1467
12.719 | -106.1472
12.7227 | -106.1496
12.7264 | -106.1521
12.73 | -106.157
12.7337 | -106.1657
12.7373 | -106.1736
12.741 | -106.1777
12.7447 | -106.1833
12.7483 | -106.1909
12.752 | -106.1989
12.7556 | -106.2094
12.7593 | -106.2214
12.763 | -106.2309
12.7666 | -106.2395
12.7703 | -106.2454
12.7739 | -106.2516
12.7776 | -106.2584
12.7813 | -106.2638
12.7849 | -106.2647
12.7886 | -106.2663
12.7922 | -106.2696
12.7959 | -106.2718
12.7996 | -106.2686
12.8032 | -106.2612
12.8069 | -106.2543
12.8105 | -106.2465
12.8142 | -106.2388
12.8179 | -106.2292
12.8215 | -106.2137
12.8252 | -106.1952
12.8288 | -106.171
12.8325 | -106.1474
12.8362 | -106.1228
12.8398 | -106.0994
12.8435 | -106.0703
12.8471 | -106.0402
12.8508 | -106.0114
12.8545 | -105.9819
12.8581 | -105.9513
12.8618 | -105.9235
12.8654 | -105.8966
12.8691 | -105.8686
12.8728 | -105.8419
12.8764 | -105.8181
12.8801 | -105.7948
12.8837 | -105.7759
12.8874 | -105.7605
12.8911 | -105.7513
12.8947 | -105.742
12.8984 | -105.7351
12.902 | -105.7292
12.9057 | -105.7272
12.9094 | -105.725
12.913 | -105.7266
12.9167 | -105.7307
12.9203 | -105.7394
12.924 | -105.7521
12.9277 | -105.7673
12.9313 | -105.7838
12.935 | -105.7969
12.9386 | -105.8102
12.9423 | -105.8253
12.946 | -105.8432
12.9496 | -105.8618
12.9533 | -105.8782
12.9569 | -105.8937
12.9606 | -105.9117
12.9643 | -105.9289
12.9679 | -105.9468
12.9716 | -105.9658
12.9752 | -105.9811
12.9789 | -105.9947
12.9826 | -106.009
12.9862 | -106.0232
12.9899 | -106.0368
12.9935 | -106.0461
12.9972 | -106.056
13.0009 | -106.0668
13.0045 | -106.0763
13.0082 | -106.0819
13.0119 | -106.0852
13.0155 | -106.0911
13.0192 | -106.0963
13.0228 | -106.099
13.0265 | -106.1007
13.0302 | -106.1046
13.0338 | -106.1022
13.0375 | -106.1011
13.0411 | -106.1024
13.0448 | -106.1039
13.0485 | -106.1055
13.0521 | -106.1087
13.0558 | -106.1132
13.0594 | -106.1189
13.0631 | -106.1249
13.0668 | -106.1307
13.0704 | -106.135
13.0741 | -106.1414
13.0777 | -106.1484
13.0814 | -106.1596
13.0851 | -106.1704
13.0887 | -106.1828
13.0924 | -106.192
13.096 | -106.2021
13.0997 | -106.2118
13.1034 | -106.2197
13.107 | -106.2264
13.1107 | -106.2317
13.1143 | -106.2387
13.118 | -106.2453
13.1217 | -106.2533
13.1253 | -106.2633
13.129 | -106.2704
13.1326 | -106.2747
13.1363 | -106.2812
13.14 | -106.285
13.1436 | -106.288
13.1473 | -106.2908
13.1509 | -106.2937
13.1546 | -106.2915
13.1583 | -106.2915
13.1619 | -106.2911
13.1656 | -106.2882
13.1692 | -106.2854
13.1729 | -106.2834
13.1766 | -106.2764
13.1802 | -106.2693
13.1839 | -106.2608
13.1875 | -106.253
13.1912 | -106.2458
13.1949 | -106.2415
13.1985 | -106.2338
13.2022 | -106.227
13.2058 | -106.2154
13.2095 | -106.2022
13.2132 | -106.1865
13.2168 | -106.1694
13.2205 | -106.1559
13.2241 | -106.1462
13.2278 | -106.1326
13.2315 | -106.1192
13.2351 | -106.1058
13.2388 | -106.0907
13.2424 | -106.0759
13.2461 | -106.0607
13.2498 | -106.0456
13.2534 | -106.0327
13.2571 | -106.0202
13.2607 | -106.0083
13.2644 | -105.9959
13.2681 | -105.9849
13.2717 | -105.9728
13.2754 | -105.965
13.279 | -105.9594
13.2827 | -105.9564
13.2864 | -105.9561
13.29 | -105.956
13.2937 | -105.9564
13.2973 | -105.9598
13.301 | -105.9645
13.3047 | -105.9687
13.3083 | -105.9753
13.312 | -105.9815
13.3156 | -105.9871
13.3193 | -105.9952
13.323 | -106.0041
13.3266 | -106.0148
13.3303 | -106.0266
13.3339 | -106.0379
13.3376 | -106.0463
13.3413 | -106.0528
13.3449 | -106.0626
13.3486 | -106.0695
13.3522 | -106.0757
13.3559 | -106.0837
13.3596 | -106.0884
13.3632 | -106.0921
13.3669 | -106.0956
13.3705 | -106.1005
13.3742 | -106.1026
13.3779 | -106.0992
13.3815 | -106.0944
13.3852 | -106.0938
13.3888 | -106.091
13.3925 | -106.088
13.3962 | -106.0884
13.3998 | -106.0904
13.4035 | -106.0926
13.4071 | -106.094
13.4108 | -106.0972
13.4145 | -106.0965
13.4181 | -106.094
13.4218 | -106.0922
13.4254 | -106.0933
13.4291 | -106.0943
13.4328 | -106.0947
13.4364 | -106.0964
13.4401 | -106.0978
13.4437 | -106.0992
13.4474 | -106.1001
13.4511 | -106.1001
13.4547 | -106.0967
13.4584 | -106.0939
13.462 | -106.0917
13.4657 | -106.0892
13.4694 | -106.0863
13.473 | -106.0825
13.4767 | -106.0799
13.4803 | -106.0765
13.484 | -106.0747
13.4877 | -106.0749
13.4913 | -106.0746
13.495 | -106.0729
13.4986 | -106.0676
13.5023 | -106.0661
13.506 | -106.0681
13.5096 | -106.0694
13.5133 | -106.0708
13.517 | -106.0732
13.5206 | -106.0726
13.5243 | -106.0742
13.5279 | -106.0791
13.5316 | -106.0826
13.5353 | -106.0886
13.5389 | -106.0948
13.5426 | -106.1001
13.5462 | -106.105
13.5499 | -106.1088
13.5536 | -106.1124
13.5572 | -106.1124
13.5609 | -106.1107
13.5645 | -106.1054
13.5682 | -106.0995
13.5719 | -106.0954
13.5755 | -106.0907
13.5792 | -106.0798
13.5828 | -106.068
13.5865 | -106.0538
13.5902 | -106.039
13.5938 | -106.0223
13.5975 | -106.0037
13.6011 | -105.9866
13.6048 | -105.9689
13.6085 | -105.951
13.6121 | -105.9357
13.6158 | -105.9192
13.6194 | -105.9003
13.6231 | -105.8833
13.6268 | -105.8709
13.6304 | -105.8601
13.6341 | -105.8522
13.6377 | -105.8452
13.6414 | -105.8411
13.6451 | -105.8411
13.6487 | -105.844
13.6524 | -105.845
13.656 | -105.8478
13.6597 | -105.8511
13.6634 | -105.854
13.667 | -105.8591
13.6707 | -105.8671
13.6743 | -105.8781
13.678 | -105.8909
13.6817 | -105.9061
13.6853 | -105.9215
13.689 | -105.94
13.6926 | -105.9585
13.6963 | -105.9732
13.7 | -105.9872
13.7036 | -106.0025
13.7073 | -106.016
13.7109 | -106.0273
13.7146 | -106.042
13.7183 | -106.0549
13.7219 | -106.0641
13.7256 | -106.0713
13.7292 | -106.083
13.7329 | -106.0933
13.7366 | -106.1016
13.7402 | -106.1108
13.7439 | -106.1191
13.7475 | -106.1264
13.7512 | -106.1353
13.7549 | -106.1456
13.7585 | -106.1572
13.7622 | -106.165
13.7658 | -106.1695
13.7695 | -106.1751
13.7732 | -106.1796
13.7768 | -106.187
13.7805 | -106.1951
13.7841 | -106.2063
13.7878 | -106.2156
13.7915 | -106.221
13.7951 | -106.2274
13.7988 | -106.237
13.8024 | -106.2449
13.8061 | -106.2552
13.8098 | -106.264
13.8134 | -106.2723
13.8171 | -106.2836
13.8207 | -106.2953
13.8244 | -106.3046
13.8281 | -106.3121
13.8317 | -106.3193
13.8354 | -106.3278
13.839 | -106.3348
13.8427 | -106.3397
13.8464 | -106.3483
13.85 | -106.3569
13.8537 | -106.3612
13.8573 | -106.3709
13.861 | -106.3802
13.8647 | -106.3821
13.8683 | -106.3808
13.872 | -106.3812
13.8756 | -106.3825
13.8793 | -106.3833
13.883 | -106.3853
13.8866 | -106.3887
13.8903 | -106.3937
13.8939 | -106.3975
13.8976 | -106.3982
13.9013 | -106.3959
13.9049 | -106.3934
13.9086 | -106.3926
13.9122 | -106.3867
13.9159 | -106.3821
13.9196 | -106.377
13.9232 | -106.369
13.9269 | -106.3578
13.9305 | -106.3471
13.9342 | -106.3356
13.9379 | -106.3234
13.9415 | -106.3123
13.9452 | -106.3021
13.9488 | -106.2914
13.9525 | -106.2798
13.9562 | -106.2666
13.9598 | -106.2487
13.9635 | -106.2265
13.9672 | -106.2078
13.9708 | -106.192
13.9745 | -106.1782
13.9781 | -106.1619
13.9818 | -106.1467
13.9855 | -106.1301
13.9891 | -106.1103
13.9928 | -106.0909
13.9964 | -106.0745
14.0001 | -106.0577
14.0038 | -106.0379
14.0074 | -106.0182
14.0111 | -105.9973
14.0147 | -105.9728
14.0184 | -105.9446
14.0221 | -105.9173
14.0257 | -105.8905
14.0294 | -105.8629
14.033 | -105.8356
14.0367 | -105.8065
14.0404 | -105.7765
14.044 | -105.7503
14.0477 | -105.7233
14.0513 | -105.6978
14.055 | -105.6754
14.0587 | -105.6533
14.0623 | -105.6354
14.066 | -105.6249
14.0696 | -105.6179
14.0733 | -105.6101
14.077 | -105.6044
14.0806 | -105.6034
14.0843 | -105.6063
14.0879 | -105.6128
14.0916 | -105.6239
14.0953 | -105.638
14.0989 | -105.6574
14.1026 | -105.6821
14.1062 | -105.7099
14.1099 | -105.7367
14.1136 | -105.7655
14.1172 | -105.7967
14.1209 | -105.831
14.1245 | -105.8685
14.1282 | -105.9055
14.1319 | -105.9405
14.1355 | -105.9748
14.1392 | -106.0123
14.1428 | -106.0473
14.1465 | -106.0793
14.1502 | -106.1114
14.1538 | -106.1448
14.1575 | -106.1787
14.1611 | -106.2143
14.1648 | -106.2522
14.1685 | -106.2842
14.1721 | -106.3067
14.1758 | -106.324
14.1794 | -106.3419
14.1831 | -106.3582
14.1868 | -106.3718
14.1904 | -106.3875
14.1941 | -106.4039
14.1977 | -106.4126
14.2014 | -106.4173
14.2051 | -106.4221
14.2087 | -106.4245
14.2124 | -106.4237
14.216 | -106.4267
14.2197 | -106.4244
14.2234 | -106.418
14.227 | -106.41
14.2307 | -106.4009
14.2343 | -106.3919
14.238 | -106.3825
14.2417 | -106.3741
14.2453 | -106.3631
14.249 | -106.3487
14.2526 | -106.3356
14.2563 | -106.3254
14.26 | -106.3164
14.2636 | -106.3068
14.2673 | -106.2978
14.2709 | -106.2892
14.2746 | -106.2824
14.2783 | -106.2768
14.2819 | -106.2742
14.2856 | -106.2715
14.2892 | -106.2697
14.2929 | -106.2702
14.2966 | -106.272
14.3002 | -106.2773
14.3039 | -106.2825
14.3075 | -106.287
14.3112 | -106.2918
14.3149 | -106.2993
14.3185 | -106.3079
14.3222 | -106.3137
14.3258 | -106.3233
14.3295 | -106.3324
14.3332 | -106.3446
14.3368 | -106.3596
14.3405 | -106.3751
14.3441 | -106.387
14.3478 | -106.3996
14.3515 | -106.4128
14.3551 | -106.4213
14.3588 | -106.4286
14.3624 | -106.437
14.3661 | -106.4439
14.3698 | -106.4475
14.3734 | -106.4519
14.3771 | -106.4549
14.3807 | -106.4581
14.3844 | -106.4592
14.3881 | -106.4576
14.3917 | -106.4544
14.3954 | -106.4514
14.399 | -106.4486
14.4027 | -106.4462
14.4064 | -106.4399
14.41 | -106.4347
14.4137 | -106.4254
14.4173 | -106.4205
14.421 | -106.4139
14.4247 | -106.4097
14.4283 | -106.4062
14.432 | -106.4029
14.4357 | -106.3989
14.4393 | -106.3931
14.443 | -106.3909
14.4466 | -106.3899
14.4503 | -106.3897
14.454 | -106.3939
14.4576 | -106.3966
14.4613 | -106.3981
14.4649 | -106.4008
14.4686 | -106.4044
14.4723 | -106.4077
14.4759 | -106.4135
14.4796 | -106.4209
14.4832 | -106.4237
14.4869 | -106.4299
14.4906 | -106.4361
14.4942 | -106.4407
14.4979 | -106.4469
14.5015 | -106.4517
14.5052 | -106.4547
14.5089 | -106.4548
14.5125 | -106.4576
14.5162 | -106.4592
14.5198 | -106.4619
14.5235 | -106.4653
14.5272 | -106.4692
14.5308 | -106.4727
14.5345 | -106.4722
14.5381 | -106.4705
14.5418 | -106.4702
14.5455 | -106.4673
14.5491 | -106.4641
14.5528 | -106.4616
14.5564 | -106.459
14.5601 | -106.4592
14.5638 | -106.4626
14.5674 | -106.4617
14.5711 | -106.462
14.5747 | -106.463
14.5784 | -106.467
14.5821 | -106.4676
14.5857 | -106.4684
14.5894 | -106.4714
14.593 | -106.4738
14.5967 | -106.4756
14.6004 | -106.4801
14.604 | -106.4867
14.6077 | -106.4931
14.6113 | -106.4996
14.615 | -106.5075
14.6187 | -106.5116
14.6223 | -106.5197
14.626 | -106.5294
14.6296 | -106.5395
14.6333 | -106.5498
14.637 | -106.5584
14.6406 | -106.569
14.6443 | -106.5805
14.6479 | -106.5926
14.6516 | -106.6036
14.6553 | -106.6135
14.6589 | -106.6217
14.6626 | -106.6304
14.6662 | -106.6393
14.6699 | -106.6513
14.6736 | -106.6603
14.6772 | -106.6659
14.6809 | -106.6713
14.6845 | -106.6757
14.6882 | -106.6774
14.6919 | -106.6816
14.6955 | -106.6846
14.6992 | -106.6853
14.7028 | -106.6867
14.7065 | -106.685
14.7102 | -106.6858
14.7138 | -106.6842
14.7175 | -106.6802
14.7211 | -106.674
14.7248 | -106.6678
14.7285 | -106.6588
14.7321 | -106.6489
14.7358 | -106.6371
14.7394 | -106.6262
14.7431 | -106.6141
14.7468 | -106.604
14.7504 | -106.5923
14.7541 | -106.579
14.7577 | -106.5651
14.7614 | -106.5494
14.7651 | -106.5348
14.7687 | -106.5206
14.7724 | -106.5054
14.776 | -106.4914
14.7797 | -106.4747
14.7834 | -106.4574
14.787 | -106.4425
14.7907 | -106.4279
14.7943 | -106.4126
14.798 | -106.3989
14.8017 | -106.3842
14.8053 | -106.3704
14.809 | -106.357
14.8126 | -106.3447
14.8163 | -106.3308
14.82 | -106.315
14.8236 | -106.3016
14.8273 | -106.2907
14.8309 | -106.2841
14.8346 | -106.2776
14.8383 | -106.2669
14.8419 | -106.2561
14.8456 | -106.2486
14.8492 | -106.2451
14.8529 | -106.2398
14.8566 | -106.2377
14.8602 | -106.2372
14.8639 | -106.2387
14.8675 | -106.2422
14.8712 | -106.2491
14.8749 | -106.2562
14.8785 | -106.2662
14.8822 | -106.2756
14.8858 | -106.2886
14.8895 | -106.3042
14.8932 | -106.3226
14.8968 | -106.3415
14.9005 | -106.3616
14.9041 | -106.3837
14.9078 | -106.4056
14.9115 | -106.4309
14.9151 | -106.4573
14.9188 | -106.4861
14.9224 | -106.5125
14.9261 | -106.5382
14.9298 | -106.5634
14.9334 | -106.5894
14.9371 | -106.6154
14.9408 | -106.6412
14.9444 | -106.6646
14.9481 | -106.6835
14.9517 | -106.7016
14.9554 | -106.7204
14.9591 | -106.7362
14.9627 | -106.7491
14.9664 | -106.7627
14.97 | -106.7719
14.9737 | -106.7786
14.9774 | -106.7786
14.981 | -106.7724
14.9847 | -106.7635
14.9883 | -106.7551
14.992 | -106.7459
14.9957 | -106.7365
14.9993 | -106.7269
15.003 | -106.7121
15.0066 | -106.6938
15.0103 | -106.6771
15.014 | -106.6582
15.0176 | -106.636
15.0213 | -106.6132
15.0249 | -106.5941
15.0286 | -106.5724
15.0323 | -106.5498
15.0359 | -106.5307
15.0396 | -106.5142
15.0432 | -106.4968
15.0469 | -106.4772
15.0506 | -106.4603
15.0542 | -106.4486
15.0579 | -106.4393
15.0615 | -106.4269
15.0652 | -106.4137
15.0689 | -106.4066
15.0725 | -106.4006
15.0762 | -106.3958
15.0798 | -106.3899
15.0835 | -106.3875
15.0872 | -106.3869
15.0908 | -106.387
15.0945 | -106.3881
15.0981 | -106.3897
15.1018 | -106.3906
15.1055 | -106.3906
15.1091 | -106.3893
15.1128 | -106.3878
15.1164 | -106.3861
15.1201 | -106.3843
15.1238 | -106.3838
15.1274 | -106.3805
15.1311 | -106.3757
15.1347 | -106.3717
15.1384 | -106.368
15.1421 | -106.3678
15.1457 | -106.3711
15.1494 | -106.3703
15.153 | -106.3669
15.1567 | -106.362
15.1604 | -106.3591
15.164 | -106.3579
15.1677 | -106.3509
15.1713 | -106.3461
15.175 | -106.3426
15.1787 | -106.3399
15.1823 | -106.3341
15.186 | -106.3305
15.1896 | -106.3283
15.1933 | -106.3243
15.197 | -106.3219
15.2006 | -106.321
15.2043 | -106.3174
15.2079 | -106.3142
15.2116 | -106.3106
15.2153 | -106.3084
15.2189 | -106.3069
15.2226 | -106.3076
15.2262 | -106.309
15.2299 | -106.3099
15.2336 | -106.3107
15.2372 | -106.3136
15.2409 | -106.3136
15.2445 | -106.3104
15.2482 | -106.3092
15.2519 | -106.3102
15.2555 | -106.3077
15.2592 | -106.3063
15.2628 | -106.3072
15.2665 | -106.3046
15.2702 | -106.3009
15.2738 | -106.2963
15.2775 | -106.295
15.2811 | -106.2946
15.2848 | -106.2919
15.2885 | -106.2883
15.2921 | -106.2864
15.2958 | -106.2863
15.2994 | -106.2852
15.3031 | -106.2851
15.3068 | -106.2861
15.3104 | -106.2893
15.3141 | -106.2908
15.3177 | -106.2933
15.3214 | -106.2962
15.3251 | -106.3022
15.3287 | -106.309
15.3324 | -106.3158
15.336 | -106.3241
15.3397 | -106.3311
15.3434 | -106.3372
15.347 | -106.3443
15.3507 | -106.3547
15.3543 | -106.3644
15.358 | -106.3779
15.3617 | -106.3906
15.3653 | -106.3994
15.369 | -106.4111
15.3726 | -106.4258
15.3763 | -106.4424
15.38 | -106.4562
15.3836 | -106.4692
15.3873 | -106.4846
15.391 | -106.498
15.3946 | -106.5115
15.3983 | -106.5249
15.4019 | -106.5408
15.4056 | -106.5536
15.4093 | -106.5632
15.4129 | -106.5778
15.4166 | -106.589
15.4202 | -106.5977
15.4239 | -106.6064
15.4276 | -106.6124
15.4312 | -106.6179
15.4349 | -106.6223
15.4385 | -106.6257
15.4422 | -106.6288
15.4459 | -106.6305
15.4495 | -106.6306
15.4532 | -106.6296
15.4568 | -106.6291
15.4605 | -106.6272
15.4642 | -106.6285
15.4678 | -106.627
15.4715 | -106.6234
15.4751 | -106.623
15.4788 | -106.621
15.4825 | -106.6191
15.4861 | -106.621
15.4898 | -106.6213
15.4934 | -106.6238
15.4971 | -106.626
15.5008 | -106.6283
15.5044 | -106.627
15.5081 | -106.6292
15.5117 | -106.6335
15.5154 | -106.6371
15.5191 | -106.6413
15.5227 | -106.6461
15.5264 | -106.6502
15.53 | -106.6512
15.5337 | -106.6561
15.5374 | -106.6581
15.541 | -106.6586
15.5447 | -106.6609
15.5483 | -106.6626
15.552 | -106.6647
15.5557 | -106.6666
15.5593 | -106.6671
15.563 | -106.6679
15.5666 | -106.6677
15.5703 | -106.6674
15.574 | -106.6636
15.5776 | -106.6586
15.5813 | -106.6516
15.5849 | -106.6445
15.5886 | -106.6388
15.5923 | -106.6307
15.5959 | -106.6235
15.5996 | -106.614
15.6032 | -106.6075
15.6069 | -106.6022
15.6106 | -106.5952
15.6142 | -106.5856
15.6179 | -106.5758
15.6215 | -106.5678
15.6252 | -106.5602
15.6289 | -106.5517
15.6325 | -106.5457
15.6362 | -106.5392
15.6398 | -106.5325
15.6435 | -106.5244
15.6472 | -106.5192
15.6508 | -106.5149
15.6545 | -106.5082
15.6581 | -106.5009
15.6618 | -106.4965
15.6655 | -106.4926
15.6691 | -106.4894
15.6728 | -106.4877
15.6764 | -106.4854
15.6801 | -106.4805
15.6838 | -106.4755
15.6874 | -106.4684
15.6911 | -106.4606
15.6947 | -106.4519
15.6984 | -106.4406
15.7021 | -106.4294
15.7057 | -106.4182
15.7094 | -106.4071
15.713 | -106.3948
15.7167 | -106.3813
15.7204 | -106.3689
15.724 | -106.3536
15.7277 | -106.3423
15.7313 | -106.3291
15.735 | -106.3142
15.7387 | -106.3
15.7423 | -106.286
15.746 | -106.2708
15.7496 | -106.2575
15.7533 | -106.2469
15.757 | -106.2327
15.7606 | -106.2165
15.7643 | -106.1992
15.7679 | -106.182
15.7716 | -106.1665
15.7753 | -106.1512
15.7789 | -106.1371
15.7826 | -106.127
15.7862 | -106.1187
15.7899 | -106.1105
15.7936 | -106.1044
15.7972 | -106.1001
15.8009 | -106.0968
15.8045 | -106.0935
15.8082 | -106.0931
15.8119 | -106.0905
15.8155 | -106.0895
15.8192 | -106.0881
15.8228 | -106.0877
15.8265 | -106.0882
15.8302 | -106.0886
15.8338 | -106.0909
15.8375 | -106.0943
15.8411 | -106.0986
15.8448 | -106.1017
15.8485 | -106.1059
15.8521 | -106.1097
15.8558 | -106.1164
15.8595 | -106.1239
15.8631 | -106.1282
15.8668 | -106.1313
15.8704 | -106.1394
15.8741 | -106.1437
15.8778 | -106.1494
15.8814 | -106.158
15.8851 | -106.1667
15.8887 | -106.1724
15.8924 | -106.1757
15.8961 | -106.1809
15.8997 | -106.1849
15.9034 | -106.1885
15.907 | -106.1936
15.9107 | -106.2019
15.9144 | -106.2093
15.918 | -106.2166
15.9217 | -106.2226
15.9253 | -106.23
15.929 | -106.2364
15.9327 | -106.243
15.9363 | -106.2465
15.94 | -106.2501
15.9436 | -106.2576
15.9473 | -106.2632
15.951 | -106.2688
15.9546 | -106.2735
15.9583 | -106.2789
15.9619 | -106.2809
15.9656 | -106.2833
15.9693 | -106.2861
15.9729 | -106.2922
15.9766 | -106.297
15.9802 | -106.302
15.9839 | -106.3071
15.9876 | -106.3123
15.9912 | -106.3179
15.9949 | -106.321
15.9985 | -106.3212
16.0022 | -106.3223
16.0059 | -106.3248
16.0095 | -106.3263
16.0132 | -106.3285
16.0168 | -106.3307
16.0205 | -106.3311
16.0242 | -106.3339
16.0278 | -106.3394
16.0315 | -106.3449
16.0351 | -106.3492
16.0388 | -106.3574
16.0425 | -106.3659
16.0461 | -106.3732
16.0498 | -106.3832
16.0534 | -106.3917
16.0571 | -106.4003
16.0608 | -106.4078
16.0644 | -106.4143
16.0681 | -106.4259
16.0717 | -106.4359
16.0754 | -106.4448
16.0791 | -106.4502
16.0827 | -106.4549
16.0864 | -106.4601
16.09 | -106.466
16.0937 | -106.4731
16.0974 | -106.4821
16.101 | -106.4876
16.1047 | -106.4944
16.1083 | -106.4987
16.112 | -106.5011
16.1157 | -106.5027
16.1193 | -106.5043
16.123 | -106.5064
16.1266 | -106.5076
16.1303 | -106.51
16.134 | -106.5119
16.1376 | -106.5086
16.1413 | -106.5045
16.1449 | -106.5047
16.1486 | -106.5048
16.1523 | -106.5038
16.1559 | -106.5051
16.1596 | -106.5062
16.1632 | -106.5062
16.1669 | -106.5074
16.1706 | -106.5112
16.1742 | -106.5114
16.1779 | -106.5135
16.1815 | -106.5196
16.1852 | -106.525
16.1889 | -106.5269
16.1925 | -106.5317
16.1962 | -106.5383
16.1998 | -106.5462
16.2035 | -106.5496
16.2072 | -106.5546
16.2108 | -106.5625
16.2145 | -106.5723
16.2181 | -106.582
16.2218 | -106.5903
16.2255 | -106.5999
16.2291 | -106.6078
16.2328 | -106.6141
16.2364 | -106.6196
16.2401 | -106.6252
16.2438 | -106.628
16.2474 | -106.6316
16.2511 | -106.6365
16.2547 | -106.6402
16.2584 | -106.6413
16.2621 | -106.6426
16.2657 | -106.6469
16.2694 | -106.6502
16.273 | -106.6523
16.2767 | -106.6534
16.2804 | -106.6541
16.284 | -106.6544
16.2877 | -106.6577
16.2913 | -106.6604
16.295 | -106.6643
16.2987 | -106.6642
16.3023 | -106.6649
16.306 | -106.6694
16.3096 | -106.6752
16.3133 | -106.6821
16.317 | -106.6892
16.3206 | -106.6947
16.3243 | -106.7016
16.328 | -106.7085
16.3316 | -106.7135
16.3353 | -106.7184
16.3389 | -106.7251
16.3426 | -106.7348
16.3463 | -106.7451
16.3499 | -106.7538
16.3536 | -106.7631
16.3572 | -106.7716
16.3609 | -106.7795
16.3646 | -106.7859
16.3682 | -106.787
16.3719 | -106.7931
16.3755 | -106.7982
16.3792 | -106.7985
16.3829 | -106.7967
16.3865 | -106.7997
16.3902 | -106.8014
16.3938 | -106.8011
16.3975 | -106.7982
16.4012 | -106.7976
16.4048 | -106.7963
16.4085 | -106.7931
16.4121 | -106.7895
16.4158 | -106.7851
16.4195 | -106.782
16.4231 | -106.781
16.4268 | -106.7792
16.4304 | -106.7746
16.4341 | -106.7688
16.4378 | -106.7654
16.4414 | -106.7616
16.4451 | -106.7567
16.4487 | -106.7516
16.4524 | -106.7502
16.4561 | -106.7502
16.4597 | -106.7479
16.4634 | -106.7486
16.467 | -106.7498
16.4707 | -106.7502
16.4744 | -106.7494
16.478 | -106.754
16.4817 | -106.7572
16.4853 | -106.7612
16.489 | -106.7648
16.4927 | -106.7675
16.4963 | -106.7691
16.5 | -106.7712
16.5036 | -106.7743
16.5073 | -106.7773
16.511 | -106.7811
16.5146 | -106.7834
16.5183 | -106.7856
16.5219 | -106.7858
16.5256 | -106.786
16.5293 | -106.7854
16.5329 | -106.7846
16.5366 | -106.7816
16.5402 | -106.779
16.5439 | -106.7763
16.5476 | -106.7718
16.5512 | -106.7668
16.5549 | -106.7612
16.5585 | -106.7565
16.5622 | -106.7523
16.5659 | -106.7484
16.5695 | -106.7426
16.5732 | -106.7368
16.5768 | -106.7295
16.5805 | -106.7225
16.5842 | -106.719
16.5878 | -106.714
16.5915 | -106.7087
16.5951 | -106.7048
16.5988 | -106.7029
16.6025 | -106.7001
16.6061 | -106.6981
16.6098 | -106.6956
16.6134 | -106.6912
16.6171 | -106.6893
16.6208 | -106.6885
16.6244 | -106.6828
16.6281 | -106.6779
16.6317 | -106.675
16.6354 | -106.6722
16.6391 | -106.6672
16.6427 | -106.6646
16.6464 | -106.6626
16.65 | -106.6599
16.6537 | -106.6568
16.6574 | -106.6557
16.661 | -106.6557
16.6647 | -106.6552
16.6683 | -106.6576
16.672 | -106.6597
16.6757 | -106.6606
16.6793 | -106.663
16.683 | -106.6652
16.6866 | -106.6704
16.6903 | -106.6752
16.694 | -106.6798
16.6976 | -106.6875
16.7013 | -106.6942
16.7049 | -106.7015
16.7086 | -106.7095
16.7123 | -106.7195
16.7159 | -106.7241
16.7196 | -106.7286
16.7232 | -106.7346
16.7269 | -106.7401
16.7306 | -106.7464
16.7342 | -106.7545
16.7379 | -106.7592
16.7415 | -106.7627
16.7452 | -106.7682
16.7489 | -106.7744
16.7525 | -106.7798
16.7562 | -106.7848
16.7598 | -106.7907
16.7635 | -106.7972
16.7672 | -106.8008
16.7708 | -106.8058
16.7745 | -106.81
16.7781 | -106.8122
16.7818 | -106.8138
16.7855 | -106.8164
16.7891 | -106.8162
16.7928 | -106.817
16.7965 | -106.8179
16.8001 | -106.8186
16.8038 | -106.8198
16.8074 | -106.8206
16.8111 | -106.8192
16.8148 | -106.8185
16.8184 | -106.8188
16.8221 | -106.8171
16.8257 | -106.8163
16.8294 | -106.8147
16.8331 | -106.8127
16.8367 | -106.8116
16.8404 | -106.8078
16.844 | -106.8061
16.8477 | -106.8065
16.8514 | -106.809
16.855 | -106.8123
16.8587 | -106.8162
16.8623 | -106.819
16.866 | -106.8195
16.8697 | -106.8219
16.8733 | -106.8235
16.877 | -106.8277
16.8806 | -106.8323
16.8843 | -106.8387
16.888 | -106.8452
16.8916 | -106.8524
16.8953 | -106.8621
16.8989 | -106.8684
16.9026 | -106.876
16.9063 | -106.8834
16.9099 | -106.8905
16.9136 | -106.8987
16.9172 | -106.9086
16.9209 | -106.9175
16.9246 | -106.9242
16.9282 | -106.9269
16.9319 | -106.9297
16.9355 | -106.9329
16.9392 | -106.9326
16.9429 | -106.9323
16.9465 | -106.9329
16.9502 | -106.9345
16.9538 | -106.9317
16.9575 | -106.93
16.9612 | -106.9315
16.9648 | -106.9288
16.9685 | -106.9226
16.9721 | -106.915
16.9758 | -106.9062
16.9795 | -106.8971
16.9831 | -106.8916
16.9868 | -106.8842
16.9904 | -106.8744
16.9941 | -106.8666
16.9978 | -106.8581
17.0014 | -106.8487
17.0051 | -106.8434
17.0087 | -106.8407
17.0124 | -106.8347
17.0161 | -106.8289
17.0197 | -106.8213
17.0234 | -106.8149
17.027 | -106.8117
17.0307 | -106.8117
17.0344 | -106.8083
17.038 | -106.8051
17.0417 | -106.8048
17.0453 | -106.8047
17.049 | -106.8058
17.0527 | -106.8088
17.0563 | -106.8088
17.06 | -106.8083
17.0636 | -106.8096
17.0673 | -106.8113
17.071 | -106.8143
17.0746 | -106.8185
17.0783 | -106.8223
17.0819 | -106.8275
17.0856 | -106.834
17.0893 | -106.8388
17.0929 | -106.8449
17.0966 | -106.8501
17.1002 | -106.8523
17.1039 | -106.8582
17.1076 | -106.8643
17.1112 | -106.8674
17.1149 | -106.8706
17.1185 | -106.8706
17.1222 | -106.8697
17.1259 | -106.8689
17.1295 | -106.8681
17.1332 | -106.8659
17.1368 | -106.8636
17.1405 | -106.86
17.1442 | -106.8571
17.1478 | -106.8528
17.1515 | -106.848
17.1551 | -106.8417
17.1588 | -106.8363
17.1625 | -106.8309
17.1661 | -106.8248
17.1698 | -106.8164
17.1734 | -106.8088
17.1771 | -106.8005
17.1808 | -106.7925
17.1844 | -106.7862
17.1881 | -106.7789
17.1917 | -106.7706
17.1954 | -106.7615
17.1991 | -106.7566
17.2027 | -106.7523
17.2064 | -106.7447
17.21 | -106.7368
17.2137 | -106.7322
17.2174 | -106.7306
17.221 | -106.7271
17.2247 | -106.7237
17.2283 | -106.7203
17.232 | -106.7199
17.2357 | -106.7226
17.2393 | -106.7246
17.243 | -106.729
17.2466 | -106.7354
17.2503 | -106.7389
17.254 | -106.743
17.2576 | -106.7504
17.2613 | -106.7603
17.2649 | -106.772
17.2686 | -106.7837
17.2723 | -106.795
17.2759 | -106.8062
17.2796 | -106.8204
17.2832 | -106.8351
17.2869 | -106.8489
17.2906 | -106.8627
17.2942 | -106.8778
17.2979 | -106.8912
17.3016 | -106.9071
17.3052 | -106.9225
17.3089 | -106.9384
17.3125 | -106.9532
17.3162 | -106.9667
17.3199 | -106.9808
17.3235 | -106.9965
17.3272 | -107.0121
17.3308 | -107.0251
17.3345 | -107.0395
17.3382 | -107.0545
17.3418 | -107.0675
17.3455 | -107.0834
17.3491 | -107.0966
17.3528 | -107.1105
17.3565 | -107.1257
17.3601 | -107.1408
17.3638 | -107.1536
17.3674 | -107.1622
17.3711 | -107.1702
17.3748 | -107.1802
17.3784 | -107.1901
17.3821 | -107.1978
17.3857 | -107.2046
17.3894 | -107.2062
17.3931 | -107.2096
17.3967 | -107.2152
17.4004 | -107.2179
17.404 | -107.2182
17.4077 | -107.2187
17.4114 | -107.2183
17.415 | -107.222
17.4187 | -107.2239
17.4223 | -107.2226
17.426 | -107.2225
17.4297 | -107.2189
17.4333 | -107.2159
17.437 | -107.2132
17.4406 | -107.2117
17.4443 | -107.2111
17.448 | -107.2098
17.4516 | -107.2093
17.4553 | -107.2088
17.4589 | -107.2078
17.4626 | -107.2068
17.4663 | -107.205
17.4699 | -107.2081
17.4736 | -107.2124
17.4772 | -107.2142
17.4809 | -107.2159
17.4846 | -107.2173
17.4882 | -107.2199
17.4919 | -107.2233
17.4955 | -107.2266
17.4992 | -107.2321
17.5029 | -107.2364
17.5065 | -107.2394
17.5102 | -107.2415
17.5138 | -107.2429
17.5175 | -107.244
17.5212 | -107.245
17.5248 | -107.2463
17.5285 | -107.2467
17.5321 | -107.2476
17.5358 | -107.2495
17.5395 | -107.2545
17.5431 | -107.2581
17.5468 | -107.2616
17.5504 | -107.2628
17.5541 | -107.2658
17.5578 | -107.2665
17.5614 | -107.2662
17.5651 | -107.2687
17.5687 | -107.2663
17.5724 | -107.2621
17.5761 | -107.2586
17.5797 | -107.2541
17.5834 | -107.2482
17.587 | -107.2418
17.5907 | -107.2373
17.5944 | -107.2314
17.598 | -107.2236
17.6017 | -107.2136
17.6053 | -107.2051
17.609 | -107.1963
17.6127 | -107.1883
17.6163 | -107.1793
17.62 | -107.17
17.6236 | -107.1595
17.6273 | -107.1468
17.631 | -107.135
17.6346 | -107.1249
17.6383 | -107.1198
17.6419 | -107.1107
17.6456 | -107.0975
17.6493 | -107.0882
17.6529 | -107.0795
17.6566 | -107.0696
17.6602 | -107.0603
17.6639 | -107.0483
17.6676 | -107.0385
17.6712 | -107.0299
17.6749 | -107.0223
17.6785 | -107.0169
17.6822 | -107.0138
17.6859 | -107.0119
17.6895 | -107.0087
17.6932 | -107.0034
17.6968 | -106.9983
17.7005 | -106.9943
17.7042 | -106.9914
17.7078 | -106.9911
17.7115 | -106.9933
17.7151 | -106.9945
17.7188 | -106.9936
17.7225 | -106.9931
17.7261 | -106.9922
17.7298 | -106.9938
17.7334 | -106.9939
17.7371 | -106.9941
17.7408 | -106.9943
17.7444 | -106.9943
17.7481 | -106.9956
17.7518 | -106.9977
17.7554 | -107.0021
17.7591 | -107.0058
17.7627 | -107.0137
17.7664 | -107.0213
17.7701 | -107.0273
17.7737 | -107.0301
17.7774 | -107.0333
17.781 | -107.0386
17.7847 | -107.0428
17.7884 | -107.0462
17.792 | -107.0499
17.7957 | -107.0559
17.7993 | -107.0622
17.803 | -107.0678
17.8067 | -107.0712
17.8103 | -107.0746
17.814 | -107.0802
17.8176 | -107.085
17.8213 | -107.0914
17.825 | -107.0983
17.8286 | -107.1043
17.8323 | -107.1085
17.8359 | -107.1119
17.8396 | -107.1127
17.8433 | -107.1136
17.8469 | -107.1161
17.8506 | -107.1198
17.8542 | -107.1234
17.8579 | -107.1237
17.8616 | -107.1229
17.8652 | -107.1216
17.8689 | -107.1238
17.8725 | -107.1255
17.8762 | -107.1258
17.8799 | -107.1258
17.8835 | -107.1265
17.8872 | -107.1298
17.8908 | -107.1313
17.8945 | -107.1355
17.8982 | -107.1422
17.9018 | -107.1461
17.9055 | -107.1492
17.9091 | -107.1543
17.9128 | -107.163
17.9165 | -107.1729
17.9201 | -107.1804
17.9238 | -107.1866
17.9274 | -107.1913
17.9311 | -107.1943
17.9348 | -107.2011
17.9384 | -107.2053
17.9421 | -107.2117
17.9457 | -107.2189
17.9494 | -107.2268
17.9531 | -107.235
17.9567 | -107.2412
17.9604 | -107.2467
17.964 | -107.2505
17.9677 | -107.2547
17.9714 | -107.2594
17.975 | -107.2628
17.9787 | -107.2629
17.9823 | -107.2653
17.986 | -107.2657
17.9897 | -107.2688
17.9933 | -107.2722
17.997 | -107.2746
18.0006 | -107.2767
18.0043 | -107.2768
18.008 | -107.2766
18.0116 | -107.2746
18.0153 | -107.2742
18.0189 | -107.2718
18.0226 | -107.2709
18.0263 | -107.2696
18.0299 | -107.268
18.0336 | -107.2655
18.0372 | -107.2625
18.0409 | -107.261
18.0446 | -107.2587
18.0482 | -107.2565
18.0519 | -107.2521
18.0555 | -107.2471
18.0592 | -107.2449
18.0629 | -107.2451
18.0665 | -107.2453
18.0702 | -107.2438
18.0738 | -107.2407
18.0775 | -107.2408
18.0812 | -107.2387
18.0848 | -107.2365
18.0885 | -107.2369
18.0921 | -107.2364
18.0958 | -107.2374
18.0995 | -107.2402
18.1031 | -107.2408
18.1068 | -107.2412
18.1104 | -107.2422
18.1141 | -107.2443
18.1178 | -107.2484
18.1214 | -107.2529
18.1251 | -107.2525
18.1287 | -107.2513
18.1324 | -107.2538
18.1361 | -107.2591
18.1397 | -107.2612
18.1434 | -107.2602
18.147 | -107.2604
18.1507 | -107.2616
18.1544 | -107.2629
18.158 | -107.2638
18.1617 | -107.2648
18.1653 | -107.2653
18.169 | -107.2663
18.1727 | -107.2684
18.1763 | -107.2711
18.18 | -107.274
18.1836 | -107.2759
18.1873 | -107.2763
18.191 | -107.2741
18.1946 | -107.2758
18.1983 | -107.2731
18.2019 | -107.2708
18.2056 | -107.2667
18.2093 | -107.2611
18.2129 | -107.2568
18.2166 | -107.2519
18.2203 | -107.2463
18.2239 | -107.2403
18.2276 | -107.2308
18.2312 | -107.2204
18.2349 | -107.2107
18.2386 | -107.197
18.2422 | -107.1824
18.2459 | -107.1697
18.2495 | -107.1539
18.2532 | -107.1366
18.2569 | -107.1198
18.2605 | -107.0991
18.2642 | -107.0764
18.2678 | -107.0527
18.2715 | -107.0307
18.2752 | -107.0105
18.2788 | -106.987
18.2825 | -106.9607
18.2861 | -106.9369
18.2898 | -106.9152
18.2935 | -106.8938
18.2971 | -106.8708
18.3008 | -106.8478
18.3044 | -106.8247
18.3081 | -106.8028
18.3118 | -106.7825
18.3154 | -106.7642
18.3191 | -106.7443
18.3227 | -106.7245
18.3264 | -106.7063
18.3301 | -106.6894
18.3337 | -106.6757
18.3374 | -106.6624
18.341 | -106.6513
18.3447 | -106.6426
18.3484 | -106.6343
18.352 | -106.6275
18.3557 | -106.6208
18.3593 | -106.6128
18.363 | -106.6075
18.3667 | -106.604
18.3703 | -106.6005
18.374 | -106.5979
18.3776 | -106.5958
18.3813 | -106.5949
18.385 | -106.5985
18.3886 | -106.6016
18.3923 | -106.6027
18.3959 | -106.6048
18.3996 | -106.6102
18.4033 | -106.6189
18.4069 | -106.6267
18.4106 | -106.6331
18.4142 | -106.6417
18.4179 | -106.6526
18.4216 | -106.6639
18.4252 | -106.675
18.4289 | -106.6835
18.4325 | -106.694
18.4362 | -106.7072
18.4399 | -106.7192
18.4435 | -106.7339
18.4472 | -106.7501
18.4508 | -106.766
18.4545 | -106.7814
18.4582 | -106.799
18.4618 | -106.816
18.4655 | -106.8323
18.4691 | -106.8467
18.4728 | -106.8607
18.4765 | -106.8753
18.4801 | -106.8902
18.4838 | -106.9057
18.4874 | -106.9209
18.4911 | -106.9362
18.4948 | -106.9492
18.4984 | -106.9618
18.5021 | -106.9764
18.5057 | -106.9893
18.5094 | -106.9984
18.5131 | -107.0069
18.5167 | -107.0168
18.5204 | -107.0267
18.524 | -107.0379
18.5277 | -107.0499
18.5314 | -107.0599
18.535 | -107.068
18.5387 | -107.0745
18.5423 | -107.0813
18.546 | -107.0896
18.5497 | -107.0967
18.5533 | -107.103
18.557 | -107.1085
18.5606 | -107.1142
18.5643 | -107.1198
18.568 | -107.1265
18.5716 | -107.1313
18.5753 | -107.1372
18.5789 | -107.1427
18.5826 | -107.1474
18.5863 | -107.1522
18.5899 | -107.1567
18.5936 | -107.1608
18.5972 | -107.1645
18.6009 | -107.1664
18.6046 | -107.1699
18.6082 | -107.1753
18.6119 | -107.1797
18.6155 | -107.1846
18.6192 | -107.1908
18.6229 | -107.2
18.6265 | -107.2097
18.6302 | -107.218
18.6338 | -107.2252
18.6375 | -107.2334
18.6412 | -107.2409
18.6448 | -107.247
18.6485 | -107.2513
18.6521 | -107.2595
18.6558 | -107.268
18.6595 | -107.274
18.6631 | -107.2773
18.6668 | -107.2771
18.6704 | -107.2767
18.6741 | -107.279
18.6778 | -107.2813
18.6814 | -107.2793
18.6851 | -107.2785
18.6887 | -107.2765
18.6924 | -107.2758
18.6961 | -107.2736
18.6997 | -107.2663
18.7034 | -107.2597
18.707 | -107.2503
18.7107 | -107.241
18.7144 | -107.2314
18.718 | -107.2211
18.7217 | -107.2117
18.7254 | -107.2029
18.729 | -107.1932
18.7327 | -107.1824
18.7363 | -107.1709
18.74 | -107.1597
18.7437 | -107.144
18.7473 | -107.1309
18.751 | -107.1195
18.7546 | -107.1057
18.7583 | -107.0913
18.762 | -107.0792
18.7656 | -107.0685
18.7693 | -107.0583
18.7729 | -107.0478
18.7766 | -107.0386
18.7803 | -107.0282
18.7839 | -107.0192
18.7876 | -107.0103
18.7912 | -107.0006
18.7949 | -106.9928
18.7986 | -106.9862
18.8022 | -106.9804
18.8059 | -106.9727
18.8095 | -106.9679
18.8132 | -106.9652
18.8169 | -106.963
18.8205 | -106.9576
18.8242 | -106.9531
18.8278 | -106.9514
18.8315 | -106.9483
18.8352 | -106.9458
18.8388 | -106.9464
18.8425 | -106.9461
18.8461 | -106.9443
18.8498 | -106.9433
18.8535 | -106.9429
18.8571 | -106.9445
18.8608 | -106.9466
18.8644 | -106.9463
18.8681 | -106.9475
18.8718 | -106.9486
18.8754 | -106.9494
18.8791 | -106.9499
18.8827 | -106.9487
18.8864 | -106.949
18.8901 | -106.9487
18.8937 | -106.948
18.8974 | -106.9468
18.901 | -106.9493
18.9047 | -106.9529
18.9084 | -106.9566
18.912 | -106.9605
18.9157 | -106.9639
18.9193 | -106.9652
18.923 | -106.9654
18.9267 | -106.9672
18.9303 | -106.9701
18.934 | -106.9748
18.9376 | -106.9811
18.9413 | -106.9855
18.945 | -106.9897
18.9486 | -106.9964
18.9523 | -107.0024
18.9559 | -107.0072
18.9596 | -107.0112
18.9633 | -107.0168
18.9669 | -107.026
18.9706 | -107.0344
18.9742 | -107.041
18.9779 | -107.0486
18.9816 | -107.0549
18.9852 | -107.057
18.9889 | -107.0601
18.9925 | -107.063
18.9962 | -107.0672
18.9999 | -107.0717
19.0035 | -107.0745
19.0072 | -107.0714
19.0108 | -107.0732
19.0145 | -107.0765
19.0182 | -107.0754
19.0218 | -107.073
19.0255 | -107.0756
19.0291 | -107.0788
19.0328 | -107.0805
19.0365 | -107.08
19.0401 | -107.0778
19.0438 | -107.0767
19.0474 | -107.0768
19.0511 | -107.0797
19.0548 | -107.0808
19.0584 | -107.0829
19.0621 | -107.088
19.0657 | -107.091
19.0694 | -107.0958
19.0731 | -107.0997
19.0767 | -107.1054
19.0804 | -107.1126
19.084 | -107.1192
19.0877 | -107.1257
19.0914 | -107.1363
19.095 | -107.1483
19.0987 | -107.1585
19.1023 | -107.1682
19.106 | -107.1782
19.1097 | -107.1857
19.1133 | -107.1954
19.117 | -107.2086
19.1206 | -107.2235
19.1243 | -107.2354
19.128 | -107.2452
19.1316 | -107.2558
19.1353 | -107.2691
19.1389 | -107.2814
19.1426 | -107.292
19.1463 | -107.3033
19.1499 | -107.3149
19.1536 | -107.3279
19.1572 | -107.3367
19.1609 | -107.3424
19.1646 | -107.3465
19.1682 | -107.3527
19.1719 | -107.3566
19.1756 | -107.3586
19.1792 | -107.3592
19.1829 | -107.3603
19.1865 | -107.3622
19.1902 | -107.3616
19.1939 | -107.359
19.1975 | -107.3555
19.2012 | -107.352
19.2048 | -107.3463
19.2085 | -107.3366
19.2122 | -107.3252
19.2158 | -107.3131
19.2195 | -107.3012
19.2231 | -107.2908
19.2268 | -107.2811
19.2305 | -107.2698
19.2341 | -107.2571
19.2378 | -107.2451
19.2414 | -107.2356
19.2451 | -107.2273
19.2488 | -107.2172
19.2524 | -107.2084
19.2561 | -107.2037
19.2597 | -107.1976
19.2634 | -107.1908
19.2671 | -107.1845
19.2707 | -107.1772
19.2744 | -107.1734
19.278 | -107.1709
19.2817 | -107.1677
19.2854 | -107.1638
19.289 | -107.1616
19.2927 | -107.16
19.2963 | -107.1582
19.3 | -107.1554
19.3037 | -107.1523
19.3073 | -107.1487
19.311 | -107.149
19.3146 | -107.1511
19.3183 | -107.1523
19.322 | -107.1524
19.3256 | -107.151
19.3293 | -107.1505
19.3329 | -107.1495
19.3366 | -107.1471
19.3403 | -107.1456
19.3439 | -107.1442
19.3476 | -107.1459
19.3512 | -107.1494
19.3549 | -107.149
19.3586 | -107.148
19.3622 | -107.1479
19.3659 | -107.1483
19.3695 | -107.1479
19.3732 | -107.1499
19.3769 | -107.1522
19.3805 | -107.1551
19.3842 | -107.1585
19.3878 | -107.1612
19.3915 | -107.1631
19.3952 | -107.1652
19.3988 | -107.168
19.4025 | -107.1716
19.4061 | -107.1738
19.4098 | -107.1759
19.4135 | -107.1792
19.4171 | -107.1834
19.4208 | -107.186
19.4244 | -107.1897
19.4281 | -107.1946
19.4318 | -107.1969
19.4354 | -107.2014
19.4391 | -107.2069
19.4427 | -107.2089
19.4464 | -107.2084
19.4501 | -107.2103
19.4537 | -107.2133
19.4574 | -107.2141
19.461 | -107.2137
19.4647 | -107.2129
19.4684 | -107.2121
19.472 | -107.2117
19.4757 | -107.2091
19.4793 | -107.2054
19.483 | -107.2026
19.4867 | -107.1992
19.4903 | -107.1926
19.494 | -107.1824
19.4976 | -107.1715
19.5013 | -107.1598
19.505 | -107.1464
19.5086 | -107.1329
19.5123 | -107.1162
19.5159 | -107.0988
19.5196 | -107.084
19.5233 | -107.0673
19.5269 | -107.0488
19.5306 | -107.0297
19.5342 | -107.0081
19.5379 | -106.9908
19.5416 | -106.9743
19.5452 | -106.9563
19.5489 | -106.938
19.5525 | -106.9161
19.5562 | -106.8947
19.5599 | -106.8774
19.5635 | -106.8603
19.5672 | -106.842
19.5708 | -106.8256
19.5745 | -106.8093
19.5782 | -106.7916
19.5818 | -106.775
19.5855 | -106.7593
19.5891 | -106.741
19.5928 | -106.7242
19.5965 | -106.7091
19.6001 | -106.6962
19.6038 | -106.6846
19.6074 | -106.6722
19.6111 | -106.6577
19.6148 | -106.6442
19.6184 | -106.6333
19.6221 | -106.6232
19.6257 | -106.6112
19.6294 | -106.5998
19.6331 | -106.5917
19.6367 | -106.5846
19.6404 | -106.5788
19.6441 | -106.573
19.6477 | -106.5671
19.6514 | -106.5596
19.655 | -106.5546
19.6587 | -106.5506
19.6624 | -106.5476
19.666 | -106.5457
19.6697 | -106.5434
19.6733 | -106.5419
19.677 | -106.5423
19.6807 | -106.5444
19.6843 | -106.548
19.688 | -106.5516
19.6916 | -106.5549
19.6953 | -106.5592
19.699 | -106.5642
19.7026 | -106.5712
19.7063 | -106.579
19.7099 | -106.5881
19.7136 | -106.5966
19.7173 | -106.6033
19.7209 | -106.6108
19.7246 | -106.6209
19.7282 | -106.629
19.7319 | -106.6378
19.7356 | -106.6477
19.7392 | -106.6572
19.7429 | -106.6675
19.7465 | -106.6776
19.7502 | -106.6858
19.7539 | -106.6935
19.7575 | -106.7023
19.7612 | -106.7106
19.7648 | -106.7194
19.7685 | -106.7294
19.7722 | -106.7399
19.7758 | -106.7466
19.7795 | -106.7544
19.7831 | -106.7627
19.7868 | -106.7683
19.7905 | -106.775
19.7941 | -106.7804
19.7978 | -106.7854
19.8014 | -106.7929
19.8051 | -106.7997
19.8088 | -106.8069
19.8124 | -106.8144
19.8161 | -106.8216
19.8197 | -106.8284
19.8234 | -106.8337
19.8271 | -106.8395
19.8307 | -106.8487
19.8344 | -106.86
19.838 | -106.8682
19.8417 | -106.873
19.8454 | -106.879
19.849 | -106.8827
19.8527 | -106.886
19.8563 | -106.889
19.86 | -106.8952
19.8637 | -106.9012
19.8673 | -106.9071
19.871 | -106.9128
19.8746 | -106.9161
19.8783 | -106.9204
19.882 | -106.9242
19.8856 | -106.9283
19.8893 | -106.9314
19.8929 | -106.9373
19.8966 | -106.9433
19.9003 | -106.9435
19.9039 | -106.9411
19.9076 | -106.9401
19.9112 | -106.9398
19.9149 | -106.9383
19.9186 | -106.9362
19.9222 | -106.9337
19.9259 | -106.9312
19.9295 | -106.9294
19.9332 | -106.9283
19.9369 | -106.9246
19.9405 | -106.9181
19.9442 | -106.9135
19.9478 | -106.9076
19.9515 | -106.9022
19.9552 | -106.8947
19.9588 | -106.8862
19.9625 | -106.8767
19.9661 | -106.8672
19.9698 | -106.8574
19.9735 | -106.8471
19.9771 | -106.839
19.9808 | -106.8307
19.9844 | -106.8197
19.9881 | -106.8082
19.9918 | -106.7998
19.9954 | -106.7919
19.9991 | -106.7854
20.0027 | -106.7784
20.0064 | -106.7676
20.0101 | -106.756
20.0137 | -106.7464
20.0174 | -106.7362
20.021 | -106.7297
20.0247 | -106.7215
20.0284 | -106.7141
20.032 | -106.7076
20.0357 | -106.7023
20.0393 | -106.6972
20.043 | -106.6942
20.0467 | -106.6904
20.0503 | -106.6877
20.054 | -106.6861
20.0576 | -106.6836
20.0613 | -106.6819
20.065 | -106.6848
20.0686 | -106.6837
20.0723 | -106.6833
20.0759 | -106.6845
20.0796 | -106.6862
20.0833 | -106.6855
20.0869 | -106.6837
20.0906 | -106.682
20.0942 | -106.6818
20.0979 | -106.6838
20.1016 | -106.686
20.1052 | -106.6861
20.1089 | -106.6881
20.1125 | -106.6901
20.1162 | -106.6923
20.1199 | -106.6926
20.1235 | -106.6951
20.1272 | -106.6991
20.1309 | -106.7019
20.1345 | -106.7081
20.1382 | -106.7137
20.1418 | -106.7187
20.1455 | -106.7218
20.1492 | -106.7237
20.1528 | -106.7242
20.1565 | -106.7282
20.1601 | -106.7311
20.1638 | -106.7314
20.1675 | -106.731
20.1711 | -106.7318
20.1748 | -106.7332
20.1784 | -106.733
20.1821 | -106.7337
20.1858 | -106.7368
20.1894 | -106.7395
20.1931 | -106.7391
20.1967 | -106.7375
20.2004 | -106.7345
20.2041 | -106.734
20.2077 | -106.7333
20.2114 | -106.7314
20.215 | -106.7295
20.2187 | -106.7273
20.2224 | -106.7259
20.226 | -106.7227
20.2297 | -106.7172
20.2333 | -106.7141
20.237 | -106.7112
20.2407 | -106.7082
20.2443 | -106.7054
20.248 | -106.7005
20.2516 | -106.6956
20.2553 | -106.692
20.259 | -106.689
20.2626 | -106.6833
20.2663 | -106.6786
20.2699 | -106.6745
20.2736 | -106.6696
20.2773 | -106.6651
20.2809 | -106.6599
20.2846 | -106.6537
20.2882 | -106.6499
20.2919 | -106.648
20.2956 | -106.6463
20.2992 | -106.6419
20.3029 | -106.638
20.3065 | -106.6317
20.3102 | -106.6252
20.3139 | -106.6183
20.3175 | -106.6139
20.3212 | -106.6083
20.3248 | -106.6011
20.3285 | -106.5925
20.3322 | -106.5855
20.3358 | -106.577
20.3395 | -106.5689
20.3431 | -106.5644
20.3468 | -106.5587
20.3505 | -106.5526
20.3541 | -106.5469
20.3578 | -106.5409
20.3614 | -106.5342
20.3651 | -106.5273
20.3688 | -106.5203
20.3724 | -106.515
20.3761 | -106.5099
20.3797 | -106.5058
20.3834 | -106.5003
20.3871 | -106.4928
20.3907 | -106.4849
20.3944 | -106.482
20.398 | -106.4791
20.4017 | -106.4758
20.4054 | -106.4711
20.409 | -106.4693
20.4127 | -106.4683
20.4163 | -106.4662
20.42 | -106.4654
20.4237 | -106.4647
20.4273 | -106.4661
20.431 | -106.4686
20.4346 | -106.4706
20.4383 | -106.4739
20.442 | -106.4784
20.4456 | -106.4808
20.4493 | -106.4865
20.4529 | -106.4914
20.4566 | -106.497
20.4603 | -106.503
20.4639 | -106.509
20.4676 | -106.5141
20.4712 | -106.5186
20.4749 | -106.524
20.4786 | -106.5313
20.4822 | -106.5403
20.4859 | -106.5481
20.4895 | -106.5564
20.4932 | -106.5623
20.4969 | -106.5675
20.5005 | -106.5735
20.5042 | -106.5798
20.5078 | -106.5861
20.5115 | -106.5928
20.5152 | -106.6012
20.5188 | -106.6101
20.5225 | -106.617
20.5261 | -106.6227
20.5298 | -106.6281
20.5335 | -106.633
20.5371 | -106.6401
20.5408 | -106.646
20.5444 | -106.6495
20.5481 | -106.6519
20.5518 | -106.6547
20.5554 | -106.656
20.5591 | -106.6588
20.5627 | -106.6624
20.5664 | -106.6627
20.5701 | -106.6644
20.5737 | -106.6671
20.5774 | -106.6661
20.5811 | -106.6638
20.5847 | -106.6593
20.5884 | -106.6536
20.592 | -106.6459
20.5957 | -106.6437
20.5994 | -106.6394
20.603 | -106.6364
20.6067 | -106.6332
20.6103 | -106.6247
20.614 | -106.6165
20.6177 | -106.6084
20.6213 | -106.6003
20.625 | -106.5913
20.6286 | -106.5825
20.6323 | -106.5745
20.636 | -106.5639
20.6396 | -106.5522
20.6433 | -106.5396
20.6469 | -106.5272
20.6506 | -106.5159
20.6543 | -106.5062
20.6579 | -106.4941
20.6616 | -106.4813
20.6652 | -106.4678
20.6689 | -106.455
20.6726 | -106.4428
20.6762 | -106.4298
20.6799 | -106.4174
20.6835 | -106.405
20.6872 | -106.3905
20.6909 | -106.3764
20.6945 | -106.3632
20.6982 | -106.3499
20.7018 | -106.3352
20.7055 | -106.3198
20.7092 | -106.3062
20.7128 | -106.2929
20.7165 | -106.2802
20.7201 | -106.2646
20.7238 | -106.2495
20.7275 | -106.239
20.7311 | -106.2298
20.7348 | -106.2196
20.7384 | -106.2084
20.7421 | -106.197
20.7458 | -106.1875
20.7494 | -106.1793
20.7531 | -106.1702
20.7567 | -106.1614
20.7604 | -106.1552
20.7641 | -106.1496
20.7677 | -106.1437
20.7714 | -106.1372
20.775 | -106.1314
20.7787 | -106.1259
20.7824 | -106.1217
20.786 | -106.1188
20.7897 | -106.1167
20.7933 | -106.1151
20.797 | -106.1146
20.8007 | -106.1148
20.8043 | -106.1145
20.808 | -106.1133
20.8116 | -106.1123
20.8153 | -106.1114
20.819 | -106.1121
20.8226 | -106.1131
20.8263 | -106.1129
20.8299 | -106.1117
20.8336 | -106.1109
20.8373 | -106.1109
20.8409 | -106.1129
20.8446 | -106.1132
20.8482 | -106.1125
20.8519 | -106.111
20.8556 | -106.1089
20.8592 | -106.1068
20.8629 | -106.1034
20.8665 | -106.1002
20.8702 | -106.0973
20.8739 | -106.0952
20.8775 | -106.0916
20.8812 | -106.0893
20.8848 | -106.0866
20.8885 | -106.0789
20.8922 | -106.0716
20.8958 | -106.0666
20.8995 | -106.0598
20.9031 | -106.053
20.9068 | -106.0484
20.9105 | -106.0427
20.9141 | -106.0372
20.9178 | -106.0313
20.9214 | -106.0261
20.9251 | -106.0211
20.9288 | -106.0148
20.9324 | -106.0085
20.9361 | -106.0028
20.9397 | -105.9976
20.9434 | -105.9921
20.9471 | -105.9881
20.9507 | -105.9835
20.9544 | -105.9786
20.958 | -105.9743
20.9617 | -105.9704
20.9654 | -105.9655
20.969 | -105.9617
20.9727 | -105.9583
20.9763 | -105.9567
20.98 | -105.9562
20.9837 | -105.9545
20.9873 | -105.9531
20.991 | -105.9528
20.9946 | -105.9507
20.9983 | -105.95
21.002 | -105.9507
21.0056 | -105.9533
21.0093 | -105.9549
21.0129 | -105.9552
21.0166 | -105.959
21.0203 | -105.9626
21.0239 | -105.9648
21.0276 | -105.9675
21.0312 | -105.9716
21.0349 | -105.9769
21.0386 | -105.9786
21.0422 | -105.9784
21.0459 | -105.9809
21.0495 | -105.9851
21.0532 | -105.989
21.0569 | -105.9944
21.0605 | -105.9984
21.0642 | -106.0003
21.0678 | -106.0027
21.0715 | -106.004
21.0752 | -106.0056
21.0788 | -106.0054
21.0825 | -106.0065
21.0862 | -106.0046
21.0898 | -106.0022
21.0935 | -105.9988
21.0971 | -105.9942
21.1008 | -105.9885
21.1045 | -105.9835
21.1081 | -105.9768
21.1118 | -105.9705
21.1154 | -105.9633
21.1191 | -105.9504
21.1228 | -105.944
21.1264 | -105.9354
21.1301 | -105.9254
21.1337 | -105.9159
21.1374 | -105.9061
21.1411 | -105.8959
21.1447 | -105.8862
21.1484 | -105.877
21.152 | -105.8654
21.1557 | -105.854
21.1594 | -105.8446
21.163 | -105.8353
21.1667 | -105.8233
21.1703 | -105.8113
21.174 | -105.7982
21.1777 | -105.7873
21.1813 | -105.7765
21.185 | -105.7648
21.1886 | -105.7561
21.1923 | -105.7478
21.196 | -105.7389
21.1996 | -105.7306
21.2033 | -105.7215
21.2069 | -105.7129
21.2106 | -105.7068
21.2143 | -105.701
21.2179 | -105.6957
21.2216 | -105.6883
21.2252 | -105.6831
21.2289 | -105.6791
21.2326 | -105.6768
21.2362 | -105.6751
21.2399 | -105.674
21.2435 | -105.6731
21.2472 | -105.6718
21.2509 | -105.6724
21.2545 | -105.673
21.2582 | -105.6729
21.2618 | -105.6744
21.2655 | -105.678
21.2692 | -105.6781
21.2728 | -105.6761
21.2765 | -105.6763
21.2801 | -105.6767
21.2838 | -105.6765
21.2875 | -105.6761
21.2911 | -105.6746
21.2948 | -105.6766
21.2984 | -105.6776
21.3021 | -105.6766
21.3058 | -105.6739
21.3094 | -105.6729
21.3131 | -105.6732
21.3167 | -105.6726
21.3204 | -105.6716
21.3241 | -105.668
21.3277 | -105.6644
21.3314 | -105.6628
21.335 | -105.6617
21.3387 | -105.6625
21.3424 | -105.6649
21.346 | -105.6633
21.3497 | -105.6597
21.3533 | -105.6568
21.357 | -105.6539
21.3607 | -105.6505
21.3643 | -105.6468
21.368 | -105.6442
21.3716 | -105.6455
21.3753 | -105.6457
21.379 | -105.6451
21.3826 | -105.643
21.3863 | -105.6434
21.3899 | -105.6423
21.3936 | -105.6401
21.3973 | -105.6395
21.4009 | -105.6378
21.4046 | -105.6374
21.4082 | -105.6381
21.4119 | -105.639
21.4156 | -105.639
21.4192 | -105.6395
21.4229 | -105.6397
21.4265 | -105.6397
21.4302 | -105.6421
21.4339 | -105.6434
21.4375 | -105.6422
21.4412 | -105.6404
21.4448 | -105.6386
21.4485 | -105.6376
21.4522 | -105.6376
21.4558 | -105.6379
21.4595 | -105.637
21.4631 | -105.6358
21.4668 | -105.6336
21.4705 | -105.6326
21.4741 | -105.6332
21.4778 | -105.6315
21.4814 | -105.628
21.4851 | -105.623
21.4888 | -105.6192
21.4924 | -105.614
21.4961 | -105.6089
21.4997 | -105.6024
21.5034 | -105.5968
21.5071 | -105.5909
21.5107 | -105.583
21.5144 | -105.5753
21.518 | -105.5674
21.5217 | -105.5596
21.5254 | -105.5541
21.529 | -105.5476
21.5327 | -105.5406
21.5363 | -105.5351
21.54 | -105.5302
21.5437 | -105.5259
21.5473 | -105.5229
21.551 | -105.5191
21.5547 | -105.5144
21.5583 | -105.5114
21.562 | -105.5118
21.5656 | -105.5118
21.5693 | -105.5113
21.573 | -105.5155
21.5766 | -105.5236
21.5803 | -105.5291
21.5839 | -105.5342
21.5876 | -105.5405
21.5913 | -105.5502
21.5949 | -105.5613
21.5986 | -105.5693
21.6022 | -105.5787
21.6059 | -105.5904
21.6096 | -105.601
21.6132 | -105.6115
21.6169 | -105.6226
21.6205 | -105.6342
21.6242 | -105.647
21.6279 | -105.6616
21.6315 | -105.6747
21.6352 | -105.6869
21.6388 | -105.6986
21.6425 | -105.7092
21.6462 | -105.7189
21.6498 | -105.7284
21.6535 | -105.7394
21.6571 | -105.749
21.6608 | -105.758
21.6645 | -105.7645
21.6681 | -105.7732
21.6718 | -105.7835
21.6754 | -105.7877
21.6791 | -105.7882
21.6828 | -105.7898
21.6864 | -105.7939
21.6901 | -105.7981
21.6937 | -105.7999
21.6974 | -105.8009
21.7011 | -105.8026
21.7047 | -105.8012
21.7084 | -105.7992
21.712 | -105.7946
21.7157 | -105.7897
21.7194 | -105.7847
21.723 | -105.7796
21.7267 | -105.7717
21.7303 | -105.7619
21.734 | -105.751
21.7377 | -105.7395
21.7413 | -105.7298
21.745 | -105.7211
21.7486 | -105.712
21.7523 | -105.7017
21.756 | -105.69
21.7596 | -105.6776
21.7633 | -105.6646
21.7669 | -105.6511
21.7706 | -105.6382
21.7743 | -105.6261
21.7779 | -105.6131
21.7816 | -105.5987
21.7852 | -105.5866
21.7889 | -105.5754
21.7926 | -105.5651
21.7962 | -105.5549
21.7999 | -105.544
21.8035 | -105.5337
21.8072 | -105.5264
21.8109 | -105.5189
21.8145 | -105.5124
21.8182 | -105.5077
21.8218 | -105.5034
21.8255 | -105.5016
21.8292 | -105.4983
21.8328 | -105.4951
21.8365 | -105.4933
21.8401 | -105.4943
21.8438 | -105.4948
21.8475 | -105.4947
21.8511 | -105.4964
21.8548 | -105.4985
21.8584 | -105.5021
21.8621 | -105.5073
21.8658 | -105.5144
21.8694 | -105.5193
21.8731 | -105.5234
21.8767 | -105.5275
21.8804 | -105.5312
21.8841 | -105.5342
21.8877 | -105.5382
21.8914 | -105.5447
21.895 | -105.553
21.8987 | -105.5593
21.9024 | -105.5643
21.906 | -105.5699
21.9097 | -105.5747
21.9133 | -105.5798
21.917 | -105.5844
21.9207 | -105.5879
21.9243 | -105.593
21.928 | -105.5959
21.9316 | -105.5976
21.9353 | -105.6006
21.939 | -105.6042
21.9426 | -105.6073
21.9463 | -105.6092
21.9499 | -105.611
21.9536 | -105.6116
21.9573 | -105.6084
21.9609 | -105.6037
21.9646 | -105.5988
21.9682 | -105.5969
21.9719 | -105.5932
21.9756 | -105.5889
21.9792 | -105.5817
21.9829 | -105.5748
21.9865 | -105.5706
21.9902 | -105.5656
21.9939 | -105.5586
21.9975 | -105.5495
22.0012 | -105.54
22.0049 | -105.5341
22.0085 | -105.5295
22.0122 | -105.5225
22.0158 | -105.5161
22.0195 | -105.5113
22.0232 | -105.5067
22.0268 | -105.4991
22.0305 | -105.4909
22.0341 | -105.4845
22.0378 | -105.4807
22.0415 | -105.4781
22.0451 | -105.4754
22.0488 | -105.4732
22.0524 | -105.4692
22.0561 | -105.4638
22.0598 | -105.4588
22.0634 | -105.4551
22.0671 | -105.4516
22.0707 | -105.4475
22.0744 | -105.4466
22.0781 | -105.4474
22.0817 | -105.4464
22.0854 | -105.4449
22.089 | -105.4439
22.0927 | -105.4437
22.0964 | -105.4444
22.1 | -105.4465
22.1037 | -105.4476
22.1073 | -105.4474
22.111 | -105.4493
22.1147 | -105.451
22.1183 | -105.452
22.122 | -105.4546
22.1256 | -105.4581
22.1293 | -105.461
22.133 | -105.4622
22.1366 | -105.4626
22.1403 | -105.4669
22.1439 | -105.4715
22.1476 | -105.4759
22.1513 | -105.4792
22.1549 | -105.4836
22.1586 | -105.4896
22.1622 | -105.4941
22.1659 | -105.4993
22.1696 | -105.5037
22.1732 | -105.5085
22.1769 | -105.5147
22.1805 | -105.5189
22.1842 | -105.5243
22.1879 | -105.5289
22.1915 | -105.5341
22.1952 | -105.5392
22.1988 | -105.544
22.2025 | -105.5492
22.2062 | -105.5535
22.2098 | -105.5557
22.2135 | -105.5557
22.2171 | -105.5587
22.2208 | -105.5644
22.2245 | -105.5703
22.2281 | -105.5741
22.2318 | -105.5781
22.2354 | -105.5812
22.2391 | -105.5817
22.2428 | -105.5828
22.2464 | -105.5842
22.2501 | -105.5848
22.2537 | -105.5852
22.2574 | -105.5851
22.2611 | -105.5855
22.2647 | -105.5857
22.2684 | -105.5855
22.272 | -105.5853
22.2757 | -105.5861
22.2794 | -105.5855
22.283 | -105.582
22.2867 | -105.5795
22.2903 | -105.5753
22.294 | -105.5726
22.2977 | -105.5665
22.3013 | -105.5612
22.305 | -105.5557
22.3086 | -105.5497
22.3123 | -105.5399
22.316 | -105.5312
22.3196 | -105.519
22.3233 | -105.5056
22.3269 | -105.4939
22.3306 | -105.4832
22.3343 | -105.4713
22.3379 | -105.4577
22.3416 | -105.4453
22.3452 | -105.4342
22.3489 | -105.4211
22.3526 | -105.4063
22.3562 | -105.3911
22.3599 | -105.3758
22.3635 | -105.3562
22.3672 | -105.3369
22.3709 | -105.3187
22.3745 | -105.3003
22.3782 | -105.2788
22.3818 | -105.2564
22.3855 | -105.2346
22.3892 | -105.2092
22.3928 | -105.18
22.3965 | -105.149
22.4001 | -105.118
22.4038 | -105.09
22.4075 | -105.0624
22.4111 | -105.0341
22.4148 | -105.0041
22.4184 | -104.9739
22.4221 | -104.9445
22.4258 | -104.9155
22.4294 | -104.886
22.4331 | -104.8564
22.4367 | -104.8257
22.4404 | -104.795
22.4441 | -104.7653
22.4477 | -104.7339
22.4514 | -104.7019
22.455 | -104.6697
22.4587 | -104.6387
22.4624 | -104.6085
22.466 | -104.5788
22.4697 | -104.5497
22.4733 | -104.5227
22.477 | -104.4976
22.4807 | -104.4745
22.4843 | -104.4533
22.488 | -104.4316
22.4916 | -104.4111
22.4953 | -104.3927
22.499 | -104.376
22.5026 | -104.3607
22.5063 | -104.3472
22.51 | -104.336
22.5136 | -104.3264
22.5173 | -104.3196
22.5209 | -104.3137
22.5246 | -104.3115
22.5283 | -104.3119
22.5319 | -104.3156
22.5356 | -104.3189
22.5392 | -104.3219
22.5429 | -104.3268
22.5466 | -104.3356
22.5502 | -104.3478
22.5539 | -104.3619
22.5575 | -104.3759
22.5612 | -104.3928
22.5649 | -104.4111
22.5685 | -104.4293
22.5722 | -104.4495
22.5758 | -104.4712
22.5795 | -104.4935
22.5832 | -104.5177
22.5868 | -104.5427
22.5905 | -104.5682
22.5941 | -104.5923
22.5978 | -104.6167
22.6015 | -104.6421
22.6051 | -104.6662
22.6088 | -104.6889
22.6124 | -104.7096
22.6161 | -104.7317
22.6198 | -104.751
22.6234 | -104.7687
22.6271 | -104.7867
22.6307 | -104.8078
22.6344 | -104.8239
22.6381 | -104.8383
22.6417 | -104.8539
22.6454 | -104.8669
22.649 | -104.8774
22.6527 | -104.8878
22.6564 | -104.8965
22.66 | -104.9047
22.6637 | -104.9103
22.6673 | -104.9165
22.671 | -104.9221
22.6747 | -104.9254
22.6783 | -104.9269
22.682 | -104.9295
22.6856 | -104.9328
22.6893 | -104.9337
22.693 | -104.9331
22.6966 | -104.9328
22.7003 | -104.9324
22.7039 | -104.9262
22.7076 | -104.9247
22.7113 | -104.9226
22.7149 | -104.9212
22.7186 | -104.9205
22.7222 | -104.9172
22.7259 | -104.9119
22.7296 | -104.9052
22.7332 | -104.8964
22.7369 | -104.8879
22.7405 | -104.8807
22.7442 | -104.8739
22.7479 | -104.8666
22.7515 | -104.8604
22.7552 | -104.8543
22.7588 | -104.848
22.7625 | -104.8401
22.7662 | -104.8315
22.7698 | -104.826
22.7735 | -104.8217
22.7771 | -104.8157
22.7808 | -104.8092
22.7845 | -104.8025
22.7881 | -104.7962
22.7918 | -104.7888
22.7954 | -104.7812
22.7991 | -104.7745
22.8028 | -104.7661
22.8064 | -104.7579
22.8101 | -104.7501
22.8137 | -104.7424
22.8174 | -104.7346
22.8211 | -104.7284
22.8247 | -104.7223
22.8284 | -104.7171
22.832 | -104.7109
22.8357 | -104.7038
22.8394 | -104.6955
22.843 | -104.687
22.8467 | -104.6763
22.8503 | -104.6678
22.854 | -104.6601
22.8577 | -104.6524
22.8613 | -104.6445
22.865 | -104.6372
22.8686 | -104.6283
22.8723 | -104.6179
22.876 | -104.6083
22.8796 | -104.601
22.8833 | -104.5926
22.8869 | -104.5836
22.8906 | -104.5722
22.8943 | -104.5608
22.8979 | -104.5509
22.9016 | -104.5406
22.9052 | -104.5291
22.9089 | -104.5192
22.9126 | -104.5091
22.9162 | -104.4987
22.9199 | -104.4891
22.9235 | -104.4797
22.9272 | -104.47
22.9309 | -104.4619
22.9345 | -104.4544
22.9382 | -104.4464
22.9418 | -104.4371
22.9455 | -104.4293
22.9492 | -104.4214
22.9528 | -104.4121
22.9565 | -104.405
22.9601 | -104.3978
22.9638 | -104.3919
22.9675 | -104.3852
22.9711 | -104.379
22.9748 | -104.3731
22.9785 | -104.3677
22.9821 | -104.362
22.9858 | -104.3575
22.9894 | -104.3525
22.9931 | -104.3475
22.9968 | -104.3439
23.0004 | -104.3421
23.0041 | -104.342
23.0077 | -104.3397
23.0114 | -104.3372
23.0151 | -104.3345
23.0187 | -104.3312
23.0224 | -104.3272
23.026 | -104.3248
23.0297 | -104.3221
23.0334 | -104.3186
23.037 | -104.3167
23.0407 | -104.3149
23.0443 | -104.3134
23.048 | -104.3124
23.0517 | -104.3095
23.0553 | -104.3078
23.059 | -104.3062
23.0626 | -104.3032
23.0663 | -104.2986
23.07 | -104.2951
23.0736 | -104.2917
23.0773 | -104.2872
23.0809 | -104.2841
23.0846 | -104.2832
23.0883 | -104.2836
23.0919 | -104.284
23.0956 | -104.2818
23.0992 | -104.2792
23.1029 | -104.2772
23.1066 | -104.2756
23.1102 | -104.2736
23.1139 | -104.2718
23.1175 | -104.2704
23.1212 | -104.2672
23.1249 | -104.2634
23.1285 | -104.2611
23.1322 | -104.2597
23.1358 | -104.2593
23.1395 | -104.2568
23.1432 | -104.2535
23.1468 | -104.2508
23.1505 | -104.2471
23.1541 | -104.2442
23.1578 | -104.2418
23.1615 | -104.2372
23.1651 | -104.233
23.1688 | -104.2299
23.1724 | -104.2267
23.1761 | -104.222
23.1798 | -104.2166
23.1834 | -104.2133
23.1871 | -104.2114
23.1907 | -104.2075
23.1944 | -104.2045
23.1981 | -104.2024
23.2017 | -104.1992
23.2054 | -104.1962
23.209 | -104.1923
23.2127 | -104.1897
23.2164 | -104.1915
23.22 | -104.1922
23.2237 | -104.1922
23.2273 | -104.1938
23.231 | -104.1959
23.2347 | -104.1971
23.2383 | -104.1952
23.242 | -104.194
23.2456 | -104.1952
23.2493 | -104.196
23.253 | -104.1964
23.2566 | -104.198
23.2603 | -104.1996
23.2639 | -104.2009
23.2676 | -104.203
23.2713 | -104.2063
23.2749 | -104.2088
23.2786 | -104.2113
23.2822 | -104.2129
23.2859 | -104.2154
23.2896 | -104.2191
23.2932 | -104.2195
23.2969 | -104.22
23.3005 | -104.2212
23.3042 | -104.2216
23.3079 | -104.2225
23.3115 | -104.2229
23.3152 | -104.2232
23.3188 | -104.223
23.3225 | -104.2228
23.3262 | -104.2223
23.3298 | -104.2207
23.3335 | -104.2176
23.3371 | -104.2158
23.3408 | -104.2138
23.3445 | -104.2105
23.3481 | -104.2088
23.3518 | -104.2085
23.3554 | -104.2062
23.3591 | -104.2048
23.3628 | -104.203
23.3664 | -104.1991
23.3701 | -104.1957
23.3737 | -104.1915
23.3774 | -104.187
23.3811 | -104.1823
23.3847 | -104.1768
23.3884 | -104.1731
23.392 | -104.1711
23.3957 | -104.1659
23.3994 | -104.1601
23.403 | -104.155
23.4067 | -104.1511
23.4103 | -104.1451
23.414 | -104.1381
23.4177 | -104.1309
23.4213 | -104.1234
23.425 | -104.1176
23.4287 | -104.1109
23.4323 | -104.1042
23.436 | -104.0974
23.4396 | -104.0895
23.4433 | -104.0819
23.447 | -104.0755
23.4506 | -104.0686
23.4543 | -104.0612
23.4579 | -104.0537
23.4616 | -104.0457
23.4653 | -104.0378
23.4689 | -104.0302
23.4726 | -104.0222
23.4762 | -104.0138
23.4799 | -104.0079
23.4836 | -104.001
23.4872 | -103.9943
23.4909 | -103.9896
23.4945 | -103.9851
23.4982 | -103.9785
23.5019 | -103.9716
23.5055 | -103.9655
23.5092 | -103.9602
23.5128 | -103.9533
23.5165 | -103.9457
23.5202 | -103.9403
23.5238 | -103.9338
23.5275 | -103.9279
23.5311 | -103.9237
23.5348 | -103.9196
23.5385 | -103.9161
23.5421 | -103.9111
23.5458 | -103.9042
23.5494 | -103.8997
23.5531 | -103.8953
23.5568 | -103.8896
23.5604 | -103.8864
23.5641 | -103.8846
23.5677 | -103.8798
23.5714 | -103.878
23.5751 | -103.8761
23.5787 | -103.8732
23.5824 | -103.8699
23.586 | -103.8683
23.5897 | -103.8663
23.5934 | -103.8634
23.597 | -103.8616
23.6007 | -103.8595
23.6043 | -103.8577
23.608 | -103.8542
23.6117 | -103.8503
23.6153 | -103.847
23.619 | -103.8428
23.6226 | -103.8392
23.6263 | -103.8365
23.63 | -103.8353
23.6336 | -103.8334
23.6373 | -103.8332
23.6409 | -103.8325
23.6446 | -103.8326
23.6483 | -103.8302
23.6519 | -103.829
23.6556 | -103.829
23.6592 | -103.829
23.6629 | -103.8291
23.6666 | -103.8285
23.6702 | -103.8271
23.6739 | -103.8269
23.6775 | -103.8267
23.6812 | -103.826
23.6849 | -103.824
23.6885 | -103.8226
23.6922 | -103.8204
23.6958 | -103.8185
23.6995 | -103.819
23.7032 | -103.8203
23.7068 | -103.821
23.7105 | -103.8206
23.7141 | -103.8223
23.7178 | -103.8225
23.7215 | -103.8208
23.7251 | -103.8196
23.7288 | -103.8226
23.7324 | -103.8237
23.7361 | -103.8226
23.7398 | -103.8217
23.7434 | -103.8238
23.7471 | -103.8251
23.7507 | -103.826
23.7544 | -103.8271
23.7581 | -103.8285
23.7617 | -103.8306
23.7654 | -103.8321
23.769 | -103.8338
23.7727 | -103.8356
23.7764 | -103.8387
23.78 | -103.8414
23.7837 | -103.8432
23.7873 | -103.8446
23.791 | -103.8463
23.7947 | -103.8487
23.7983 | -103.8512
23.802 | -103.8534
23.8056 | -103.8564
23.8093 | -103.8596
23.813 | -103.8608
23.8166 | -103.8622
23.8203 | -103.8648
23.8239 | -103.8686
23.8276 | -103.8719
23.8313 | -103.8746
23.8349 | -103.8776
23.8386 | -103.8807
23.8422 | -103.8842
23.8459 | -103.8868
23.8496 | -103.8905
23.8532 | -103.8945
23.8569 | -103.8973
23.8605 | -103.9007
23.8642 | -103.9029
23.8679 | -103.9058
23.8715 | -103.9086
23.8752 | -103.9117
23.8788 | -103.916
23.8825 | -103.9191
23.8862 | -103.9217
23.8898 | -103.925
23.8935 | -103.9288
23.8971 | -103.932
23.9008 | -103.9352
23.9045 | -103.9396
23.9081 | -103.9428
23.9118 | -103.9454
23.9154 | -103.9456
23.9191 | -103.9464
23.9228 | -103.9467
23.9264 | -103.9481
23.9301 | -103.9501
23.9338 | -103.9522
23.9374 | -103.952
23.9411 | -103.9504
23.9447 | -103.9506
23.9484 | -103.9522
23.9521 | -103.9542
23.9557 | -103.9537
23.9594 | -103.9521
23.963 | -103.9496
23.9667 | -103.9479
23.9704 | -103.9458
23.974 | -103.9418
23.9777 | -103.9385
23.9813 | -103.9332
23.985 | -103.9291
23.9887 | -103.9247
23.9923 | -103.919
23.996 | -103.913
23.9996 | -103.9073
24.0033 | -103.9029
24.007 | -103.8972
24.0106 | -103.8911
24.0143 | -103.8853
24.0179 | -103.8787
24.0216 | -103.8727
24.0253 | -103.8655
24.0289 | -103.8584
24.0326 | -103.8517
24.0362 | -103.8439
24.0399 | -103.8331
24.0436 | -103.8237
24.0472 | -103.8166
24.0509 | -103.8075
24.0545 | -103.7994
24.0582 | -103.7919
24.0619 | -103.7857
24.0655 | -103.7781
24.0692 | -103.7708
24.0728 | -103.765
24.0765 | -103.7582
24.0802 | -103.7507
24.0838 | -103.7449
24.0875 | -103.7397
24.0911 | -103.7348
24.0948 | -103.7312
24.0985 | -103.7263
24.1021 | -103.7208
24.1058 | -103.7162
24.1094 | -103.7118
24.1131 | -103.7085
24.1168 | -103.7068
24.1204 | -103.7062
24.1241 | -103.7045
24.1277 | -103.7042
24.1314 | -103.7028
24.1351 | -103.7022
24.1387 | -103.703
24.1424 | -103.7041
24.146 | -103.7036
24.1497 | -103.7021
24.1534 | -103.7032
24.157 | -103.7036
24.1607 | -103.7042
24.1643 | -103.7042
24.168 | -103.7053
24.1717 | -103.706
24.1753 | -103.7075
24.179 | -103.7099
24.1826 | -103.7112
24.1863 | -103.7122
24.19 | -103.713
24.1936 | -103.7136
24.1973 | -103.7137
24.2009 | -103.7147
24.2046 | -103.7155
24.2083 | -103.7138
24.2119 | -103.7122
24.2156 | -103.7107
24.2192 | -103.7084
24.2229 | -103.7065
24.2266 | -103.7024
24.2302 | -103.6963
24.2339 | -103.6908
24.2375 | -103.6828
24.2412 | -103.6741
24.2449 | -103.6657
24.2485 | -103.6571
24.2522 | -103.6477
24.2558 | -103.6358
24.2595 | -103.6236
24.2632 | -103.6111
24.2668 | -103.5951
24.2705 | -103.5694
24.2741 | -103.5217
24.2778 | -103.4679
24.2815 | -103.4143
24.2851 | -103.3454
24.2888 | -103.227
24.2924 | -103.1758
24.2961 | -103.1247
24.2998 | -103.0712
24.3034 | -103.0166
24.3071 | -102.9601
24.3107 | -102.899
24.3144 | -102.8359
24.3181 | -102.7698
24.3217 | -102.7025
24.3254 | -102.6354
24.329 | -102.5662
24.3327 | -102.4939
24.3364 | -102.4203
24.34 | -102.3451
24.3437 | -102.2688
24.3473 | -102.1927
24.351 | -102.1166
24.3547 | -102.0395
24.3583 | -101.9616
24.362 | -101.8836
24.3656 | -101.8063
24.3693 | -101.7295
24.373 | -101.6538
24.3766 | -101.5795
24.3803 | -101.5067
24.3839 | -101.4371
24.3876 | -101.3688
24.3913 | -101.3021
24.3949 | -101.2389
24.3986 | -101.1788
24.4023 | -101.1222
24.4059 | -101.0672
24.4096 | -101.0166
24.4132 | -100.9702
24.4169 | -100.9272
24.4206 | -100.8889
24.4242 | -100.8552
24.4279 | -100.8254
24.4315 | -100.8002
24.4352 | -100.7793
24.4389 | -100.7632
24.4425 | -100.7534
24.4462 | -100.7484
24.4498 | -100.7482
24.4535 | -100.7531
24.4572 | -100.7613
24.4608 | -100.773
24.4645 | -100.7898
24.4681 | -100.8127
24.4718 | -100.8406
24.4755 | -100.8737
24.4791 | -100.9102
24.4828 | -100.9511
24.4864 | -100.9962
24.4901 | -101.0455
24.4938 | -101.0986
24.4974 | -101.1541
24.5011 | -101.2131
24.5047 | -101.2749
24.5084 | -101.3385
24.5121 | -101.4049
24.5157 | -101.4735
24.5194 | -101.5434
24.523 | -101.6148
24.5267 | -101.6865
24.5304 | -101.7585
24.534 | -101.8314
24.5377 | -101.9045
24.5413 | -101.9776
24.545 | -102.051
24.5487 | -102.1233
24.5523 | -102.1932
24.556 | -102.2605
24.5596 | -102.3251
24.5633 | -102.3816
24.567 | -102.4292
24.5706 | -102.4638
24.5743 | -102.4864
24.5779 | -102.4846
24.5816 | -102.4616
24.5853 | -102.4274
24.5889 | -102.3896
24.5926 | -102.3431
24.5962 | -102.2997
24.5999 | -102.2627
24.6036 | -102.2233
24.6072 | -102.1776
24.6109 | -102.1264
24.6145 | -102.0734
24.6182 | -102.0242
24.6219 | -101.9684
24.6255 | -101.8999
24.6292 | -101.83
24.6328 | -101.753
24.6365 | -101.6536
24.6402 | -101.5436
24.6438 | -101.4296
24.6475 | -101.3155
24.6511 | -101.1974
24.6548 | -101.0781
24.6585 | -100.9589
24.6621 | -100.8394
24.6658 | -100.7209
24.6694 | -100.6029
24.6731 | -100.4856
24.6768 | -100.3709
24.6804 | -100.2589
24.6841 | -100.1495
24.6877 | -100.0425
24.6914 | -99.9381
24.6951 | -99.8367
24.6987 | -99.7401
24.7024 | -99.6482
24.706 | -99.5613
24.7097 | -99.4797
24.7134 | -99.4027
24.717 | -99.3307
24.7207 | -99.2647
24.7243 | -99.2039
24.728 | -99.1498
24.7317 | -99.1017
24.7353 | -99.0602
24.739 | -99.0251
24.7426 | -98.9962
24.7463 | -98.9737
24.75 | -98.9578
24.7536 | -98.9482
24.7573 | -98.9454
24.7609 | -98.949
24.7646 | -98.9596
24.7683 | -98.9771
24.7719 | -99.0006
24.7756 | -99.0304
24.7792 | -99.0665
24.7829 | -99.1093
24.7866 | -99.1586
24.7902 | -99.2132
24.7939 | -99.2739
24.7975 | -99.3405
24.8012 | -99.412
24.8049 | -99.4889
24.8085 | -99.5717
24.8122 | -99.6589
24.8158 | -99.7684
24.8195 | -99.8594
24.8232 | -99.954
24.8268 | -100.0516
24.8305 | -100.1528
24.8341 | -100.2565
24.8378 | -100.3628
24.8415 | -100.4715
24.8451 | -100.5816
24.8488 | -100.6932
24.8525 | -100.8047
24.8561 | -100.916
24.8598 | -101.0291
24.8634 | -101.1417
24.8671 | -101.2525
24.8708 | -101.3627
24.8744 | -101.4708
24.8781 | -101.578
24.8817 | -101.6824
24.8854 | -101.7833
24.8891 | -101.8822
24.8927 | -101.9767
24.8964 | -102.066
24.9 | -102.1532
24.9037 | -102.2379
24.9074 | -102.3188
24.911 | -102.3964
24.9147 | -102.4754
24.9183 | -102.5558
24.922 | -102.6392
24.9257 | -102.7217
24.9293 | -102.8127
24.933 | -102.9093
24.9366 | -102.9959
24.9403 | -103.0665
24.944 | -103.1285
24.9476 | -103.1697
24.9513 | -103.1905
24.9549 | -103.2026
24.9586 | -103.2092
24.9623 | -103.2133
24.9659 | -103.2182
24.9696 | -103.2224
24.9732 | -103.2266
24.9769 | -103.2331
24.9806 | -103.238
24.9842 | -103.2422
24.9879 | -103.246
24.9915 | -103.2481
24.9952 | -103.2498
24.9989 | -103.2506
25.0025 | -103.253
25.0062 | -103.2544
25.0098 | -103.2558
25.0135 | -103.2565
25.0172 | -103.2565
25.0208 | -103.2546
25.0245 | -103.2528
25.0281 | -103.2508
25.0318 | -103.2487
25.0355 | -103.2475
25.0391 | -103.2472
25.0428 | -103.2487
25.0464 | -103.2508
25.0501 | -103.2512
25.0538 | -103.251
25.0574 | -103.2519
25.0611 | -103.2506
25.0647 | -103.2482
25.0684 | -103.2477
25.0721 | -103.2463
25.0757 | -103.2451
25.0794 | -103.242
25.083 | -103.2395
25.0867 | -103.2366
25.0904 | -103.2319
25.094 | -103.2287
25.0977 | -103.2276
25.1013 | -103.2264
25.105 | -103.2273
25.1087 | -103.2261
25.1123 | -103.2242
25.116 | -103.2223
25.1196 | -103.2189
25.1233 | -103.2159
25.127 | -103.2148
25.1306 | -103.214
25.1343 | -103.2137
25.1379 | -103.2117
25.1416 | -103.2099
25.1453 | -103.2078
25.1489 | -103.2041
25.1526 | -103.2002
25.1562 | -103.1983
25.1599 | -103.1964
25.1636 | -103.1944
25.1672 | -103.1922
25.1709 | -103.1903
25.1745 | -103.1894
25.1782 | -103.1873
25.1819 | -103.1857
25.1855 | -103.1851
25.1892 | -103.1826
25.1928 | -103.1804
25.1965 | -103.1792
25.2002 | -103.1794
25.2038 | -103.1792
25.2075 | -103.1787
25.2111 | -103.1794
25.2148 | -103.177
25.2185 | -103.1752
25.2221 | -103.1757
25.2258 | -103.1751
25.2294 | -103.1745
25.2331 | -103.1734
25.2368 | -103.1744
25.2404 | -103.1744
25.2441 | -103.1731
25.2477 | -103.1732
25.2514 | -103.173
25.2551 | -103.1734
25.2587 | -103.175
25.2624 | -103.1759
25.266 | -103.1758
25.2697 | -103.1742
25.2734 | -103.1736
25.277 | -103.1729
25.2807 | -103.1726
25.2843 | -103.174
25.288 | -103.1753
25.2917 | -103.1775
25.2953 | -103.1787
25.299 | -103.18
25.3026 | -103.1814
25.3063 | -103.1836
25.31 | -103.1841
25.3136 | -103.1857
25.3173 | -103.1888
25.3209 | -103.1919
25.3246 | -103.1951
25.3283 | -103.1956
25.3319 | -103.1965
25.3356 | -103.2005
25.3393 | -103.2051
25.3429 | -103.211
25.3466 | -103.2157
25.3502 | -103.2215
25.3539 | -103.2254
25.3576 | -103.2289
25.3612 | -103.2343
25.3649 | -103.2405
25.3685 | -103.2463
25.3722 | -103.2535
25.3759 | -103.2605
25.3795 | -103.268
25.3832 | -103.2748
25.3868 | -103.2803
25.3905 | -103.2856
25.3942 | -103.2905
25.3978 | -103.2947
25.4015 | -103.2984
25.4051 | -103.3018
25.4088 | -103.3047
25.4125 | -103.3085
25.4161 | -103.3096
25.4198 | -103.3107
25.4234 | -103.3148
25.4271 | -103.3187
25.4308 | -103.3207
25.4344 | -103.3221
25.4381 | -103.3235
25.4417 | -103.3229
25.4454 | -103.3219
25.4491 | -103.3236
25.4527 | -103.3225
25.4564 | -103.3193
25.46 | -103.3171
25.4637 | -103.312
25.4674 | -103.3085
25.471 | -103.3061
25.4747 | -103.304
25.4783 | -103.3033
25.482 | -103.3022
25.4857 | -103.299
25.4893 | -103.2953
25.493 | -103.2914
25.4966 | -103.2881
25.5003 | -103.2867
25.504 | -103.2844
25.5076 | -103.2838
25.5113 | -103.2847
25.5149 | -103.2843
25.5186 | -103.2841
25.5223 | -103.2839
25.5259 | -103.2841
25.5296 | -103.2823
25.5332 | -103.2812
25.5369 | -103.282
25.5406 | -103.2838
25.5442 | -103.2827
25.5479 | -103.284
25.5515 | -103.2858
25.5552 | -103.2891
25.5589 | -103.2922
25.5625 | -103.2961
25.5662 | -103.3007
25.5698 | -103.3042
25.5735 | -103.3097
25.5772 | -103.315
25.5808 | -103.3203
25.5845 | -103.3264
25.5881 | -103.3327
25.5918 | -103.3389
25.5955 | -103.3441
25.5991 | -103.3507
25.6028 | -103.3579
25.6064 | -103.3645
25.6101 | -103.3718
25.6138 | -103.3788
25.6174 | -103.3854
25.6211 | -103.3902
25.6247 | -103.3952
25.6284 | -103.4025
25.6321 | -103.4077
25.6357 | -103.4115
25.6394 | -103.4162
25.643 | -103.4209
25.6467 | -103.4241
25.6504 | -103.4282
25.654 | -103.4323
25.6577 | -103.4345
25.6613 | -103.4358
25.665 | -103.4365
25.6687 | -103.4375
25.6723 | -103.4372
25.676 | -103.437
25.6796 | -103.4359
25.6833 | -103.4356
25.687 | -103.4368
25.6906 | -103.4369
25.6943 | -103.4351
25.6979 | -103.433
25.7016 | -103.4298
25.7053 | -103.4248
25.7089 | -103.4206
25.7126 | -103.4164
25.7162 | -103.4123
25.7199 | -103.4077
25.7236 | -103.4006
25.7272 | -103.3935
25.7309 | -103.3881
25.7345 | -103.3818
25.7382 | -103.3771
25.7419 | -103.3726
25.7455 | -103.3671
25.7492 | -103.3602
25.7528 | -103.3543
25.7565 | -103.3485
25.7602 | -103.3429
25.7638 | -103.3381
25.7675 | -103.3322
25.7711 | -103.327
25.7748 | -103.3205
25.7785 | -103.3152
25.7821 | -103.3115
25.7858 | -103.3062
25.7894 | -103.3002
25.7931 | -103.2957
25.7968 | -103.2913
25.8004 | -103.2857
25.8041 | -103.2817
25.8077 | -103.2775
25.8114 | -103.2717
25.8151 | -103.2677
25.8187 | -103.2634
25.8224 | -103.257
25.8261 | -103.2541
25.8297 | -103.251
25.8334 | -103.248
25.837 | -103.2443
25.8407 | -103.2398
25.8444 | -103.2363
25.848 | -103.2333
25.8517 | -103.2317
25.8553 | -103.2301
25.859 | -103.229
25.8627 | -103.2265
25.8663 | -103.2233
25.87 | -103.2232
25.8736 | -103.2274
25.8773 | -103.2283
25.881 | -103.227
25.8846 | -103.2249
25.8883 | -103.2241
25.8919 | -103.2243
25.8956 | -103.2263
25.8993 | -103.227
25.9029 | -103.2267
25.9066 | -103.2248
25.9102 | -103.223
25.9139 | -103.2217
25.9176 | -103.2213
25.9212 | -103.2198
25.9249 | -103.2194
25.9285 | -103.2206
25.9322 | -103.222
25.9359 | -103.2239
25.9395 | -103.225
25.9432 | -103.2263
25.9468 | -103.2274
25.9505 | -103.2296
25.9542 | -103.2322
25.9578 | -103.2347
25.9615 | -103.2377
25.9651 | -103.2395
25.9688 | -103.2416
25.9725 | -103.2439
25.9761 | -103.2471
25.9798 | -103.2482
25.9834 | -103.2498
25.9871 | -103.2554
25.9908 | -103.2622
25.9944 | -103.2667
25.9981 | -103.2705
26.0017 | -103.275
26.0054 | -103.279
26.0091 | -103.2834
26.0127 | -103.2881
26.0164 | -103.2905
26.02 | -103.2935
26.0237 | -103.2978
26.0274 | -103.303
26.031 | -103.3067
26.0347 | -103.3094
26.0383 | -103.3142
26.042 | -103.3185
26.0457 | -103.3216
26.0493 | -103.3235
26.053 | -103.3261
26.0566 | -103.3294
26.0603 | -103.3328
26.064 | -103.3363
26.0676 | -103.3408
26.0713 | -103.3466
26.0749 | -103.3496
26.0786 | -103.3514
26.0823 | -103.3551
26.0859 | -103.3593
26.0896 | -103.3641
26.0932 | -103.3696
26.0969 | -103.3744
26.1006 | -103.3783
26.1042 | -103.3806
26.1079 | -103.3845
26.1115 | -103.3886
26.1152 | -103.3935
26.1189 | -103.3986
26.1225 | -103.4011
26.1262 | -103.4032
26.1298 | -103.4059
26.1335 | -103.4094
26.1372 | -103.4127
26.1408 | -103.4159
26.1445 | -103.4195
26.1481 | -103.4227
26.1518 | -103.4248
26.1555 | -103.4274
26.1591 | -103.4296
26.1628 | -103.4323
26.1664 | -103.4337
26.1701 | -103.4361
26.1738 | -103.4396
26.1774 | -103.4441
26.1811 | -103.4477
26.1847 | -103.4519
26.1884 | -103.4559
26.1921 | -103.4593
26.1957 | -103.4632
26.1994 | -103.4656
26.203 | -103.4681
26.2067 | -103.4715
26.2104 | -103.4726
26.214 | -103.4734
26.2177 | -103.4749
26.2213 | -103.4765
26.225 | -103.4745
26.2287 | -103.4745
26.2323 | -103.4755
26.236 | -103.4705
26.2396 | -103.4641
26.2433 | -103.4597
26.247 | -103.4568
26.2506 | -103.4566
26.2543 | -103.4565
26.2579 | -103.4541
26.2616 | -103.4491
26.2653 | -103.4455
26.2689 | -103.4435
26.2726 | -103.4402
26.2763 | -103.4373
26.2799 | -103.435
26.2836 | -103.4295
26.2872 | -103.42
26.2909 | -103.4116
26.2946 | -103.4013
26.2982 | -103.3897
26.3019 | -103.3768
26.3055 | -103.3648
26.3092 | -103.3545
26.3129 | -103.3428
26.3165 | -103.3299
26.3202 | -103.3154
26.3238 | -103.3011
26.3275 | -103.2872
26.3312 | -103.273
26.3348 | -103.2595
26.3385 | -103.2451
26.3421 | -103.231
26.3458 | -103.2153
26.3495 | -103.2016
26.3531 | -103.1881
26.3568 | -103.1724
26.3604 | -103.1572
26.3641 | -103.1439
26.3678 | -103.1307
26.3714 | -103.1176
26.3751 | -103.1054
26.3787 | -103.0917
26.3824 | -103.0785
26.3861 | -103.068
26.3897 | -103.0585
26.3934 | -103.0488
26.397 | -103.0398
26.4007 | -103.0394
26.4044 | -103.0354
26.408 | -103.0315
26.4117 | -103.0259
26.4153 | -103.0206
26.419 | -103.0169
26.4227 | -103.0151
26.4263 | -103.013
26.43 | -103.0121
26.4336 | -103.0111
26.4373 | -103.0109
26.441 | -103.0129
26.4446 | -103.0165
26.4483 | -103.0215
26.4519 | -103.025
26.4556 | -103.0291
26.4593 | -103.0344
26.4629 | -103.0401
26.4666 | -103.0448
26.4702 | -103.0491
26.4739 | -103.0552
26.4776 | -103.061
26.4812 | -103.0651
26.4849 | -103.0696
26.4885 | -103.077
26.4922 | -103.0825
26.4959 | -103.0896
26.4995 | -103.098
26.5032 | -103.1046
26.5068 | -103.1114
26.5105 | -103.1175
26.5142 | -103.1231
26.5178 | -103.129
26.5215 | -103.1333
26.5251 | -103.1375
26.5288 | -103.1407
26.5325 | -103.1438
26.5361 | -103.1462
26.5398 | -103.1507
26.5434 | -103.1528
26.5471 | -103.1532
26.5508 | -103.1527
26.5544 | -103.152
26.5581 | -103.1508
26.5617 | -103.1516
26.5654 | -103.1517
26.5691 | -103.1505
26.5727 | -103.1502
26.5764 | -103.1506
26.58 | -103.1489
26.5837 | -103.1485
26.5874 | -103.148
26.591 | -103.1454
26.5947 | -103.1436
26.5983 | -103.1448
26.602 | -103.1449
26.6057 | -103.1455
26.6093 | -103.1485
26.613 | -103.1508
26.6166 | -103.1509
26.6203 | -103.1494
26.624 | -103.1485
26.6276 | -103.147
26.6313 | -103.1466
26.6349 | -103.1485
26.6386 | -103.1495
26.6423 | -103.1489
26.6459 | -103.1484
26.6496 | -103.1481
26.6532 | -103.1484
26.6569 | -103.1507
26.6606 | -103.1555
26.6642 | -103.1593
26.6679 | -103.1625
26.6715 | -103.1648
26.6752 | -103.1691
26.6789 | -103.1711
26.6825 | -103.1694
26.6862 | -103.1625
26.6898 | -103.1509
26.6935 | -103.1341
26.6972 | -103.1141
26.7008 | -103.0968
26.7045 | -103.0809
26.7081 | -103.068
26.7118 | -103.0567
26.7155 | -103.0457
26.7191 | -103.0336
26.7228 | -103.021
26.7264 | -103.0065
26.7301 | -102.9923
26.7338 | -102.9754
26.7374 | -102.955
26.7411 | -102.9345
26.7447 | -102.9122
26.7484 | -102.8888
26.7521 | -102.8665
26.7557 | -102.8424
26.7594 | -102.8154
26.7631 | -102.7869
26.7667 | -102.7576
26.7704 | -102.7263
26.774 | -102.6941
26.7777 | -102.66
26.7814 | -102.6247
26.785 | -102.5893
26.7887 | -102.5543
26.7923 | -102.518
26.796 | -102.4816
26.7997 | -102.4448
26.8033 | -102.4084
26.807 | -102.3701
26.8106 | -102.332
26.8143 | -102.2951
26.818 | -102.2579
26.8216 | -102.2211
26.8253 | -102.1862
26.8289 | -102.1534
26.8326 | -102.1211
26.8363 | -102.0898
26.8399 | -102.0603
26.8436 | -102.0338
26.8472 | -102.0093
26.8509 | -101.986
26.8546 | -101.9659
26.8582 | -101.9482
26.8619 | -101.9332
26.8655 | -101.919
26.8692 | -101.9074
26.8729 | -101.8985
26.8765 | -101.8915
26.8802 | -101.8869
26.8838 | -101.8853
26.8875 | -101.8872
26.8912 | -101.8917
26.8948 | -101.8987
26.8985 | -101.9094
26.9021 | -101.9231
26.9058 | -101.9394
26.9095 | -101.9571
26.9131 | -101.9782
26.9168 | -102.001
26.9204 | -102.0264
26.9241 | -102.0555
26.9278 | -102.1048
26.9314 | -102.1377
26.9351 | -102.1718
26.9387 | -102.2067
26.9424 | -102.2431
26.9461 | -102.2805
26.9497 | -102.3183
26.9534 | -102.3582
26.957 | -102.3988
26.9607 | -102.442
26.9644 | -102.4846
26.968 | -102.5284
26.9717 | -102.5731
26.9753 | -102.6172
26.979 | -102.6602
26.9827 | -102.7031
26.9863 | -102.7457
26.99 | -102.7872
26.9936 | -102.8282
26.9973 | -102.8684
27.001 | -102.9096
27.0046 | -102.9481
27.0083 | -102.984
27.0119 | -103.0186
27.0156 | -103.0529
27.0193 | -103.0864
27.0229 | -103.1193
27.0266 | -103.1508
27.0302 | -103.1794
27.0339 | -103.2062
27.0376 | -103.2315
27.0412 | -103.2546
27.0449 | -103.2735
27.0485 | -103.2925
27.0522 | -103.3103
27.0559 | -103.3263
27.0595 | -103.3425
27.0632 | -103.359
27.0668 | -103.3793
27.0705 | -103.4021
27.0742 | -103.426
27.0778 | -103.4463
27.0815 | -103.4596
27.0851 | -103.4679
27.0888 | -103.4693
27.0925 | -103.4659
27.0961 | -103.462
27.0998 | -103.4587
27.1034 | -103.4524
27.1071 | -103.445
27.1108 | -103.4376
27.1144 | -103.4305
27.1181 | -103.424
27.1217 | -103.4164
27.1254 | -103.41
27.1291 | -103.4039
27.1327 | -103.3961
27.1364 | -103.3876
27.14 | -103.3798
27.1437 | -103.3722
27.1474 | -103.3637
27.151 | -103.3559
27.1547 | -103.3494
27.1583 | -103.3419
27.162 | -103.3336
27.1657 | -103.3253
27.1693 | -103.3165
27.173 | -103.3105
27.1766 | -103.3027
27.1803 | -103.2951
27.184 | -103.2888
27.1876 | -103.2846
27.1913 | -103.2799
27.1949 | -103.2736
27.1986 | -103.2689
27.2023 | -103.2643
27.2059 | -103.2583
27.2096 | -103.2528
27.2132 | -103.2475
27.2169 | -103.2425
27.2206 | -103.2397
27.2242 | -103.2363
27.2279 | -103.2335
27.2316 | -103.2305
27.2352 | -103.2273
27.2389 | -103.2259
27.2425 | -103.2256
27.2462 | -103.2245
27.2499 | -103.2238
27.2535 | -103.2231
27.2572 | -103.2223
27.2608 | -103.2208
27.2645 | -103.2193
27.2682 | -103.2169
27.2718 | -103.2174
27.2755 | -103.2171
27.2791 | -103.2166
27.2828 | -103.2164
27.2865 | -103.2148
27.2901 | -103.2158
27.2938 | -103.2174
27.2974 | -103.2181
27.3011 | -103.2183
27.3048 | -103.217
27.3084 | -103.2162
27.3121 | -103.2162
27.3157 | -103.2163
27.3194 | -103.2166
27.3231 | -103.219
27.3267 | -103.2209
27.3304 | -103.2228
27.334 | -103.2256
27.3377 | -103.2274
27.3414 | -103.2288
27.345 | -103.2295
27.3487 | -103.2311
27.3523 | -103.2323
27.356 | -103.2326
27.3597 | -103.2335
27.3633 | -103.2346
27.367 | -103.2366
27.3706 | -103.2388
27.3743 | -103.2418
27.378 | -103.2456
27.3816 | -103.2491
27.3853 | -103.2498
27.3889 | -103.2513
27.3926 | -103.2568
27.3963 | -103.2615
27.3999 | -103.2634
27.4036 | -103.2668
27.4072 | -103.27
27.4109 | -103.2729
27.4146 | -103.2772
27.4182 | -103.282
27.4219 | -103.2847
27.4255 | -103.2849
27.4292 | -103.287
27.4329 | -103.2891
27.4365 | -103.2914
27.4402 | -103.2933
27.4438 | -103.2955
27.4475 | -103.298
27.4512 | -103.3011
27.4548 | -103.3041
27.4585 | -103.3055
27.4621 | -103.3075
27.4658 | -103.3096
27.4695 | -103.3111
27.4731 | -103.3123
27.4768 | -103.3124
27.4804 | -103.3135
27.4841 | -103.3141
27.4878 | -103.3135
27.4914 | -103.3138
27.4951 | -103.3151
27.4987 | -103.3153
27.5024 | -103.3146
27.5061 | -103.3162
27.5097 | -103.3171
27.5134 | -103.3174
27.517 | -103.3143
27.5207 | -103.3107
27.5244 | -103.3092
27.528 | -103.3071
27.5317 | -103.3035
27.5353 | -103.3
27.539 | -103.2972
27.5427 | -103.2931
27.5463 | -103.2892
27.55 | -103.2843
27.5536 | -103.2794
27.5573 | -103.275
27.561 | -103.269
27.5646 | -103.2648
27.5683 | -103.2602
27.5719 | -103.2542
27.5756 | -103.2474
27.5793 | -103.2414
27.5829 | -103.2365
27.5866 | -103.2306
27.5902 | -103.2244
27.5939 | -103.2183
27.5976 | -103.2125
27.6012 | -103.2069
27.6049 | -103.2009
27.6085 | -103.1934
27.6122 | -103.1877
27.6159 | -103.1823
27.6195 | -103.1777
27.6232 | -103.173
27.6268 | -103.1668
27.6305 | -103.1605
27.6342 | -103.1565
27.6378 | -103.1533
27.6415 | -103.1486
27.6451 | -103.1446
27.6488 | -103.1421
27.6525 | -103.1386
27.6561 | -103.1369
27.6598 | -103.134
27.6634 | -103.1313
27.6671 | -103.1291
27.6708 | -103.1265
27.6744 | -103.1266
27.6781 | -103.124
27.6818 | -103.1208
27.6854 | -103.1187
27.6891 | -103.1173
27.6927 | -103.118
27.6964 | -103.1174
27.7001 | -103.1166
27.7037 | -103.1165
27.7074 | -103.1169
27.711 | -103.1169
27.7147 | -103.1165
27.7184 | -103.1167
27.722 | -103.1176
27.7257 | -103.1179
27.7293 | -103.117
27.733 | -103.1164
27.7367 | -103.1177
27.7403 | -103.1192
27.744 | -103.1204
27.7476 | -103.1189
27.7513 | -103.1181
27.755 | -103.1168
27.7586 | -103.1158
27.7623 | -103.1146
27.7659 | -103.1146
27.7696 | -103.1151
27.7733 | -103.1165
27.7769 | -103.1166
27.7806 | -103.1147
27.7842 | -103.1133
27.7879 | -103.1136
27.7916 | -103.112
27.7952 | -103.1112
27.7989 | -103.1106
27.8025 | -103.1102
27.8062 | -103.1098
27.8099 | -103.1097
27.8135 | -103.1103
27.8172 | -103.1086
27.8208 | -103.1071
27.8245 | -103.1056
27.8282 | -103.1051
27.8318 | -103.1035
27.8355 | -103.1007
27.8391 | -103.0976
27.8428 | -103.0946
27.8465 | -103.0936
27.8501 | -103.0914
27.8538 | -103.0888
27.8574 | -103.0868
27.8611 | -103.0848
27.8648 | -103.0825
27.8684 | -103.0814
27.8721 | -103.0803
27.8757 | -103.0775
27.8794 | -103.075
27.8831 | -103.0733
27.8867 | -103.0719
27.8904 | -103.069
27.894 | -103.0658
27.8977 | -103.0611
27.9014 | -103.057
27.905 | -103.053
27.9087 | -103.0498
27.9123 | -103.0459
27.916 | -103.0416
27.9197 | -103.0369
27.9233 | -103.0309
27.927 | -103.0223
27.9306 | -103.0121
27.9343 | -103.0045
27.938 | -102.9991
27.9416 | -102.9943
27.9453 | -102.9874
27.9489 | -102.9814
27.9526 | -102.9758
27.9563 | -102.9682
27.9599 | -102.9596
27.9636 | -102.9502
27.9672 | -102.9418
27.9709 | -102.9351
27.9746 | -102.9285
27.9782 | -102.9194
27.9819 | -102.9097
27.9855 | -102.8982
27.9892 | -102.8891
27.9929 | -102.8798
27.9965 | -102.8679
28.0002 | -102.854