## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3d | 26 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 2
Hood | 02-91C/P | 14
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1001 | hPa |  | PRC PRO 01
Humidity | 31.4 | %/Hr
Internal temperature of calibration bench | 23.26 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1009 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 23.4 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.04 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.32 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.95 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3186 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Analog acquisition system | Analog Hood
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 05/03/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 13/03/2018
Technical Manager | Anthony HUE | Approval |  | 20/03/2018
CEO | Robert GLEMOT | Approval |  | 20/03/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.2716 | 57.6033
0.012209 | -35.8156 | 38.5043
0.018313 | -34.7728 | 28.5321
0.024417 | -34.3084 | 22.1782
0.030521 | -34.1966 | 18.3469
0.036626 | -34.0522 | 15.2594
0.04273 | -33.9674 | 12.9552
0.048834 | -33.916 | 11.3693
0.054938 | -33.8478 | 10.1519
0.061043 | -33.7755 | 9.0229
0.067147 | -33.7636 | 8.1631
0.073251 | -33.7641 | 7.4125
0.079355 | -33.7814 | 6.8445
0.08546 | -33.7972 | 6.3701
0.091564 | -33.7926 | 5.9856
0.097668 | -33.7612 | 5.7307
0.10377 | -33.7172 | 5.3931
0.10988 | -33.6877 | 5.0357
0.11598 | -33.6776 | 4.6778
0.12209 | -33.6911 | 4.3951
0.12819 | -33.7197 | 4.158
0.13429 | -33.7478 | 3.9155
0.1404 | -33.7627 | 3.6408
0.1465 | -33.7762 | 3.381
0.15261 | -33.7823 | 3.1397
0.15871 | -33.7791 | 2.9411
0.16482 | -33.7663 | 2.7812
0.17092 | -33.7451 | 2.7158
0.17702 | -33.7287 | 2.617
0.18313 | -33.7136 | 2.5334
0.18923 | -33.7026 | 2.429
0.19534 | -33.6962 | 2.3266
0.20144 | -33.6949 | 2.2145
0.20754 | -33.6969 | 2.1369
0.21365 | -33.6988 | 2.0588
0.21975 | -33.7025 | 1.9898
0.22586 | -33.7097 | 1.9206
0.23196 | -33.7176 | 1.8452
0.23807 | -33.7257 | 1.766
0.24417 | -33.728 | 1.6745
0.25027 | -33.7314 | 1.5908
0.25638 | -33.7315 | 1.5006
0.26248 | -33.7299 | 1.4149
0.26859 | -33.7278 | 1.3327
0.27469 | -33.7271 | 1.2474
0.2808 | -33.7324 | 1.1818
0.2869 | -33.7364 | 1.1125
0.293 | -33.7413 | 1.0529
0.29911 | -33.747 | 1.0103
0.30521 | -33.7523 | 0.96926
0.31132 | -33.7557 | 0.94978
0.31742 | -33.7594 | 0.92274
0.32353 | -33.7622 | 0.90035
0.32963 | -33.7646 | 0.8813
0.33573 | -33.766 | 0.86506
0.34184 | -33.768 | 0.84637
0.34794 | -33.7706 | 0.81148
0.35405 | -33.7734 | 0.78085
0.36015 | -33.7764 | 0.74709
0.36626 | -33.7783 | 0.70983
0.37236 | -33.7804 | 0.66077
0.37846 | -33.7823 | 0.60446
0.38457 | -33.7815 | 0.54527
0.39067 | -33.782 | 0.49492
0.39678 | -33.7806 | 0.43856
0.40288 | -33.7786 | 0.38487
0.40899 | -33.7753 | 0.33204
0.41509 | -33.7714 | 0.2789
0.42119 | -33.7671 | 0.23014
0.4273 | -33.7638 | 0.18013
0.4334 | -33.7608 | 0.12953
0.43951 | -33.7582 | 0.079785
0.44561 | -33.757 | 0.033033
0.45172 | -33.7584 | -0.0055705
0.45782 | -33.7583 | -0.048434
0.46392 | -33.7595 | -0.087945
0.47003 | -33.7611 | -0.12403
0.47613 | -33.763 | -0.15732
0.48224 | -33.765 | -0.18893
0.48834 | -33.7671 | -0.21237
0.49445 | -33.7689 | -0.23904
0.50055 | -33.7711 | -0.26333
0.50665 | -33.7726 | -0.28406
0.51276 | -33.7736 | -0.30609
0.51886 | -33.7746 | -0.32606
0.52497 | -33.7737 | -0.33594
0.53107 | -33.7741 | -0.34929
0.53717 | -33.7744 | -0.3635
0.54328 | -33.7742 | -0.37206
0.54938 | -33.7742 | -0.38316
0.55549 | -33.7738 | -0.39492
0.56159 | -33.7738 | -0.40802
0.5677 | -33.7734 | -0.42073
0.5738 | -33.7731 | -0.43599
0.5799 | -33.7732 | -0.45019
0.58601 | -33.7736 | -0.46816
0.59211 | -33.7746 | -0.49123
0.59822 | -33.7753 | -0.5091
0.60432 | -33.7763 | -0.5316
0.61043 | -33.7772 | -0.55358
0.61653 | -33.7783 | -0.57594
0.62263 | -33.7793 | -0.59977
0.62874 | -33.7806 | -0.62121
0.63484 | -33.7816 | -0.64321
0.64095 | -33.7827 | -0.6652
0.64705 | -33.7838 | -0.6862
0.65316 | -33.7847 | -0.70693
0.65926 | -33.7856 | -0.73054
0.66536 | -33.7851 | -0.75035
0.67147 | -33.7853 | -0.76818
0.67757 | -33.7857 | -0.78745
0.68368 | -33.7855 | -0.80882
0.68978 | -33.7855 | -0.8282
0.69589 | -33.7851 | -0.84805
0.70199 | -33.7843 | -0.87094
0.70809 | -33.7837 | -0.89223
0.7142 | -33.7832 | -0.91267
0.7203 | -33.7823 | -0.93335
0.72641 | -33.7815 | -0.95523
0.73251 | -33.7812 | -0.97696
0.73862 | -33.7811 | -1.0019
0.74472 | -33.7813 | -1.0224
0.75082 | -33.7808 | -1.0443
0.75693 | -33.7808 | -1.0685
0.76303 | -33.7807 | -1.0918
0.76914 | -33.7812 | -1.1192
0.77524 | -33.7814 | -1.1447
0.78135 | -33.7817 | -1.1688
0.78745 | -33.7823 | -1.1942
0.79355 | -33.7827 | -1.2193
0.79966 | -33.7836 | -1.2455
0.80576 | -33.7847 | -1.2759
0.81187 | -33.7851 | -1.3056
0.81797 | -33.7857 | -1.3364
0.82408 | -33.7863 | -1.364
0.83018 | -33.7868 | -1.3953
0.83628 | -33.7876 | -1.4258
0.84239 | -33.7875 | -1.4543
0.84849 | -33.7876 | -1.4828
0.8546 | -33.788 | -1.5109
0.8607 | -33.7881 | -1.5405
0.86681 | -33.7882 | -1.5693
0.87291 | -33.7879 | -1.5955
0.87901 | -33.7874 | -1.6188
0.88512 | -33.7873 | -1.6426
0.89122 | -33.7872 | -1.6668
0.89733 | -33.787 | -1.6872
0.90343 | -33.7867 | -1.7095
0.90953 | -33.7869 | -1.7257
0.91564 | -33.7864 | -1.7431
0.92174 | -33.7861 | -1.7589
0.92785 | -33.7856 | -1.7741
0.93395 | -33.7856 | -1.7875
0.94006 | -33.7857 | -1.8025
0.94616 | -33.7855 | -1.8094
0.95226 | -33.7856 | -1.8225
0.95837 | -33.7857 | -1.8358
0.96447 | -33.7861 | -1.8478
0.97058 | -33.7861 | -1.8595
0.97668 | -33.7859 | -1.8695
0.98279 | -33.786 | -1.8755
0.98889 | -33.7862 | -1.8825
0.99499 | -33.7862 | -1.893
1.0011 | -33.7865 | -1.9
1.0072 | -33.7864 | -1.9061
1.0133 | -33.7863 | -1.9122
1.0194 | -33.7863 | -1.92
1.0255 | -33.7861 | -1.929
1.0316 | -33.7855 | -1.9355
1.0377 | -33.7851 | -1.9437
1.0438 | -33.7844 | -1.9505
1.0499 | -33.783 | -1.9592
1.056 | -33.782 | -1.9665
1.0621 | -33.7809 | -1.9746
1.0682 | -33.7805 | -1.9818
1.0743 | -33.7795 | -1.9918
1.0805 | -33.7789 | -2.0009
1.0866 | -33.7775 | -2.0135
1.0927 | -33.7762 | -2.0232
1.0988 | -33.7755 | -2.0356
1.1049 | -33.7742 | -2.0452
1.111 | -33.7727 | -2.0581
1.1171 | -33.7715 | -2.0693
1.1232 | -33.7703 | -2.0849
1.1293 | -33.7692 | -2.0985
1.1354 | -33.7684 | -2.1125
1.1415 | -33.7673 | -2.1281
1.1476 | -33.7661 | -2.1436
1.1537 | -33.7648 | -2.1582
1.1598 | -33.7647 | -2.1775
1.1659 | -33.7642 | -2.1911
1.172 | -33.7635 | -2.2057
1.1781 | -33.7629 | -2.2211
1.1842 | -33.7624 | -2.2379
1.1903 | -33.763 | -2.2541
1.1964 | -33.7629 | -2.2698
1.2025 | -33.7629 | -2.2863
1.2086 | -33.7633 | -2.3031
1.2147 | -33.7635 | -2.3193
1.2209 | -33.7642 | -2.334
1.227 | -33.7659 | -2.3501
1.2331 | -33.7668 | -2.3655
1.2392 | -33.7677 | -2.3824
1.2453 | -33.7688 | -2.3958
1.2514 | -33.77 | -2.4093
1.2575 | -33.7714 | -2.4246
1.2636 | -33.7736 | -2.4373
1.2697 | -33.7751 | -2.4503
1.2758 | -33.7766 | -2.4656
1.2819 | -33.7782 | -2.4802
1.288 | -33.7798 | -2.4951
1.2941 | -33.7817 | -2.5071
1.3002 | -33.7835 | -2.517
1.3063 | -33.7851 | -2.5284
1.3124 | -33.7867 | -2.5399
1.3185 | -33.7886 | -2.5519
1.3246 | -33.7903 | -2.564
1.3307 | -33.7922 | -2.5749
1.3368 | -33.7939 | -2.584
1.3429 | -33.7954 | -2.5943
1.349 | -33.7968 | -2.6044
1.3551 | -33.7984 | -2.6166
1.3613 | -33.7999 | -2.6261
1.3674 | -33.8011 | -2.6361
1.3735 | -33.8023 | -2.6457
1.3796 | -33.8034 | -2.6548
1.3857 | -33.8044 | -2.6635
1.3918 | -33.8054 | -2.6747
1.3979 | -33.8066 | -2.6847
1.404 | -33.8073 | -2.6962
1.4101 | -33.8082 | -2.7068
1.4162 | -33.809 | -2.7169
1.4223 | -33.8097 | -2.7287
1.4284 | -33.8105 | -2.7386
1.4345 | -33.8111 | -2.7504
1.4406 | -33.8116 | -2.7631
1.4467 | -33.8122 | -2.7741
1.4528 | -33.8129 | -2.7849
1.4589 | -33.8137 | -2.797
1.465 | -33.8139 | -2.8064
1.4711 | -33.8148 | -2.817
1.4772 | -33.8153 | -2.8306
1.4833 | -33.8159 | -2.8442
1.4894 | -33.8166 | -2.8571
1.4955 | -33.8173 | -2.8701
1.5016 | -33.818 | -2.8835
1.5078 | -33.8184 | -2.8994
1.5139 | -33.8191 | -2.9134
1.52 | -33.8196 | -2.9279
1.5261 | -33.82 | -2.941
1.5322 | -33.8202 | -2.9551
1.5383 | -33.8205 | -2.9687
1.5444 | -33.8211 | -2.9833
1.5505 | -33.8215 | -2.9979
1.5566 | -33.822 | -3.0121
1.5627 | -33.8226 | -3.0269
1.5688 | -33.8234 | -3.0414
1.5749 | -33.8236 | -3.0552
1.581 | -33.824 | -3.0729
1.5871 | -33.8245 | -3.0897
1.5932 | -33.8249 | -3.104
1.5993 | -33.8252 | -3.1218
1.6054 | -33.8256 | -3.138
1.6115 | -33.826 | -3.1559
1.6176 | -33.8263 | -3.1733
1.6237 | -33.8265 | -3.189
1.6298 | -33.8268 | -3.2055
1.6359 | -33.8272 | -3.2213
1.642 | -33.8277 | -3.2369
1.6482 | -33.828 | -3.2549
1.6543 | -33.8285 | -3.2728
1.6604 | -33.8289 | -3.2885
1.6665 | -33.8289 | -3.3046
1.6726 | -33.8292 | -3.3219
1.6787 | -33.8297 | -3.3399
1.6848 | -33.8296 | -3.3574
1.6909 | -33.8299 | -3.3746
1.697 | -33.8302 | -3.3918
1.7031 | -33.8302 | -3.4086
1.7092 | -33.8302 | -3.4247
1.7153 | -33.8304 | -3.4419
1.7214 | -33.8301 | -3.4584
1.7275 | -33.8299 | -3.4753
1.7336 | -33.8297 | -3.4925
1.7397 | -33.8298 | -3.5089
1.7458 | -33.8298 | -3.5265
1.7519 | -33.8297 | -3.5437
1.758 | -33.8296 | -3.5597
1.7641 | -33.8297 | -3.5754
1.7702 | -33.8295 | -3.5922
1.7763 | -33.8295 | -3.6072
1.7824 | -33.8294 | -3.6244
1.7885 | -33.8292 | -3.6397
1.7947 | -33.8293 | -3.6557
1.8008 | -33.8292 | -3.6699
1.8069 | -33.8292 | -3.6848
1.813 | -33.8292 | -3.7003
1.8191 | -33.829 | -3.7158
1.8252 | -33.8288 | -3.731
1.8313 | -33.8288 | -3.7466
1.8374 | -33.8289 | -3.7614
1.8435 | -33.8289 | -3.775
1.8496 | -33.8287 | -3.7907
1.8557 | -33.8287 | -3.8067
1.8618 | -33.8287 | -3.8199
1.8679 | -33.8287 | -3.8342
1.874 | -33.8287 | -3.8487
1.8801 | -33.8286 | -3.863
1.8862 | -33.8288 | -3.8772
1.8923 | -33.8287 | -3.8908
1.8984 | -33.8285 | -3.903
1.9045 | -33.8285 | -3.9156
1.9106 | -33.8282 | -3.9289
1.9167 | -33.828 | -3.9416
1.9228 | -33.828 | -3.954
1.9289 | -33.8277 | -3.9646
1.9351 | -33.8278 | -3.9753
1.9412 | -33.8279 | -3.9872
1.9473 | -33.828 | -4.0005
1.9534 | -33.828 | -4.0131
1.9595 | -33.8278 | -4.026
1.9656 | -33.8274 | -4.0382
1.9717 | -33.8273 | -4.0503
1.9778 | -33.8274 | -4.0632
1.9839 | -33.8274 | -4.076
1.99 | -33.8275 | -4.0874
1.9961 | -33.8273 | -4.0993
2.0022 | -33.8273 | -4.1108
2.0083 | -33.827 | -4.1223
2.0144 | -33.827 | -4.1337
2.0205 | -33.8271 | -4.1469
2.0266 | -33.8272 | -4.1598
2.0327 | -33.8273 | -4.1724
2.0388 | -33.8276 | -4.1873
2.0449 | -33.8277 | -4.1991
2.051 | -33.8278 | -4.2106
2.0571 | -33.8277 | -4.2233
2.0632 | -33.8276 | -4.2353
2.0693 | -33.8277 | -4.2479
2.0754 | -33.8279 | -4.2609
2.0816 | -33.8281 | -4.2742
2.0877 | -33.8281 | -4.2867
2.0938 | -33.8283 | -4.3004
2.0999 | -33.8283 | -4.3133
2.106 | -33.8286 | -4.3265
2.1121 | -33.8286 | -4.3401
2.1182 | -33.8288 | -4.3516
2.1243 | -33.829 | -4.3635
2.1304 | -33.8291 | -4.3752
2.1365 | -33.8295 | -4.388
2.1426 | -33.8296 | -4.4025
2.1487 | -33.8301 | -4.4154
2.1548 | -33.8305 | -4.4296
2.1609 | -33.831 | -4.4429
2.167 | -33.8315 | -4.4566
2.1731 | -33.8318 | -4.4699
2.1792 | -33.8325 | -4.4824
2.1853 | -33.833 | -4.4964
2.1914 | -33.8334 | -4.5099
2.1975 | -33.834 | -4.5233
2.2036 | -33.8344 | -4.5377
2.2097 | -33.8348 | -4.5509
2.2158 | -33.8357 | -4.5626
2.222 | -33.8363 | -4.5741
2.2281 | -33.8366 | -4.5863
2.2342 | -33.8371 | -4.6005
2.2403 | -33.8375 | -4.6121
2.2464 | -33.8383 | -4.6245
2.2525 | -33.8387 | -4.6367
2.2586 | -33.8395 | -4.648
2.2647 | -33.8403 | -4.6588
2.2708 | -33.8409 | -4.6718
2.2769 | -33.8417 | -4.6846
2.283 | -33.8422 | -4.6955
2.2891 | -33.8427 | -4.7073
2.2952 | -33.8432 | -4.7186
2.3013 | -33.8439 | -4.7303
2.3074 | -33.8445 | -4.7414
2.3135 | -33.845 | -4.7541
2.3196 | -33.8456 | -4.7643
2.3257 | -33.8459 | -4.7763
2.3318 | -33.8463 | -4.7869
2.3379 | -33.8467 | -4.7996
2.344 | -33.8471 | -4.8119
2.3501 | -33.8475 | -4.8241
2.3562 | -33.8478 | -4.8341
2.3623 | -33.8481 | -4.8444
2.3685 | -33.8484 | -4.855
2.3746 | -33.8486 | -4.867
2.3807 | -33.8491 | -4.8776
2.3868 | -33.8493 | -4.8882
2.3929 | -33.8497 | -4.8991
2.399 | -33.85 | -4.9088
2.4051 | -33.8505 | -4.9199
2.4112 | -33.8508 | -4.9305
2.4173 | -33.8511 | -4.941
2.4234 | -33.8511 | -4.9513
2.4295 | -33.8514 | -4.9615
2.4356 | -33.8516 | -4.9725
2.4417 | -33.8517 | -4.9841
2.4478 | -33.8518 | -4.9952
2.4539 | -33.8521 | -5.0048
2.46 | -33.8518 | -5.0161
2.4661 | -33.8519 | -5.027
2.4722 | -33.8522 | -5.0374
2.4783 | -33.8523 | -5.0475
2.4844 | -33.8523 | -5.0578
2.4905 | -33.8523 | -5.0686
2.4966 | -33.8521 | -5.0791
2.5027 | -33.8521 | -5.0889
2.5089 | -33.8521 | -5.0991
2.515 | -33.8519 | -5.11
2.5211 | -33.8519 | -5.12
2.5272 | -33.8517 | -5.1308
2.5333 | -33.8515 | -5.1419
2.5394 | -33.8516 | -5.1526
2.5455 | -33.8514 | -5.1636
2.5516 | -33.8514 | -5.174
2.5577 | -33.8514 | -5.1856
2.5638 | -33.8513 | -5.1975
2.5699 | -33.8511 | -5.2093
2.576 | -33.8509 | -5.2213
2.5821 | -33.8508 | -5.2328
2.5882 | -33.8507 | -5.244
2.5943 | -33.8506 | -5.2554
2.6004 | -33.8505 | -5.2669
2.6065 | -33.8504 | -5.2771
2.6126 | -33.8503 | -5.2878
2.6187 | -33.8503 | -5.2983
2.6248 | -33.8502 | -5.309
2.6309 | -33.8501 | -5.3217
2.637 | -33.8502 | -5.3344
2.6431 | -33.8502 | -5.3455
2.6492 | -33.8501 | -5.3571
2.6554 | -33.8502 | -5.3689
2.6615 | -33.85 | -5.3804
2.6676 | -33.85 | -5.3924
2.6737 | -33.8502 | -5.4047
2.6798 | -33.8503 | -5.4163
2.6859 | -33.8503 | -5.4283
2.692 | -33.8503 | -5.4397
2.6981 | -33.8503 | -5.4523
2.7042 | -33.8504 | -5.4657
2.7103 | -33.8505 | -5.4774
2.7164 | -33.8507 | -5.4896
2.7225 | -33.8509 | -5.5024
2.7286 | -33.851 | -5.515
2.7347 | -33.8511 | -5.5265
2.7408 | -33.8514 | -5.5374
2.7469 | -33.8514 | -5.5503
2.753 | -33.8517 | -5.5624
2.7591 | -33.8519 | -5.5758
2.7652 | -33.8519 | -5.5887
2.7713 | -33.852 | -5.6002
2.7774 | -33.8524 | -5.613
2.7835 | -33.8526 | -5.6252
2.7896 | -33.8528 | -5.6362
2.7958 | -33.8532 | -5.6481
2.8019 | -33.8537 | -5.6595
2.808 | -33.854 | -5.6727
2.8141 | -33.8544 | -5.6869
2.8202 | -33.8548 | -5.7004
2.8263 | -33.8553 | -5.7128
2.8324 | -33.8556 | -5.7266
2.8385 | -33.856 | -5.74
2.8446 | -33.8566 | -5.7531
2.8507 | -33.857 | -5.7666
2.8568 | -33.8575 | -5.7792
2.8629 | -33.8579 | -5.7919
2.869 | -33.8584 | -5.8047
2.8751 | -33.8588 | -5.818
2.8812 | -33.8595 | -5.8325
2.8873 | -33.8598 | -5.8453
2.8934 | -33.8602 | -5.8586
2.8995 | -33.8606 | -5.8719
2.9056 | -33.8609 | -5.8841
2.9117 | -33.8614 | -5.8971
2.9178 | -33.8621 | -5.9111
2.9239 | -33.8625 | -5.9256
2.93 | -33.863 | -5.9403
2.9361 | -33.8636 | -5.9532
2.9423 | -33.8642 | -5.9665
2.9484 | -33.8647 | -5.9812
2.9545 | -33.8654 | -5.9957
2.9606 | -33.8659 | -6.0102
2.9667 | -33.8665 | -6.0235
2.9728 | -33.8669 | -6.037
2.9789 | -33.8674 | -6.0511
2.985 | -33.8678 | -6.066
2.9911 | -33.8683 | -6.0798
2.9972 | -33.8689 | -6.0938
3.0033 | -33.8695 | -6.1081
3.0094 | -33.87 | -6.1224
3.0155 | -33.8706 | -6.1363
3.0216 | -33.8711 | -6.1509
3.0277 | -33.8716 | -6.1658
3.0338 | -33.8722 | -6.1806
3.0399 | -33.8728 | -6.1955
3.046 | -33.8732 | -6.2098
3.0521 | -33.8738 | -6.2236
3.0582 | -33.8741 | -6.2386
3.0643 | -33.8747 | -6.2529
3.0704 | -33.8749 | -6.2675
3.0765 | -33.8753 | -6.2811
3.0827 | -33.8757 | -6.2954
3.0888 | -33.8761 | -6.3095
3.0949 | -33.8764 | -6.3237
3.101 | -33.8769 | -6.3389
3.1071 | -33.8772 | -6.3542
3.1132 | -33.8777 | -6.3694
3.1193 | -33.8782 | -6.3836
3.1254 | -33.8786 | -6.3971
3.1315 | -33.8791 | -6.4126
3.1376 | -33.8793 | -6.4267
3.1437 | -33.8797 | -6.4414
3.1498 | -33.8801 | -6.4568
3.1559 | -33.8805 | -6.4712
3.162 | -33.8806 | -6.4864
3.1681 | -33.8807 | -6.5017
3.1742 | -33.8811 | -6.5166
3.1803 | -33.8815 | -6.531
3.1864 | -33.8819 | -6.5446
3.1925 | -33.8822 | -6.5594
3.1986 | -33.8824 | -6.5735
3.2047 | -33.8825 | -6.5886
3.2108 | -33.8828 | -6.6017
3.2169 | -33.8832 | -6.6166
3.223 | -33.8833 | -6.6312
3.2292 | -33.8833 | -6.6462
3.2353 | -33.8833 | -6.6573
3.2414 | -33.8835 | -6.6716
3.2475 | -33.8835 | -6.684
3.2536 | -33.8839 | -6.6989
3.2597 | -33.8841 | -6.7115
3.2658 | -33.8843 | -6.7236
3.2719 | -33.8845 | -6.7361
3.278 | -33.8846 | -6.7495
3.2841 | -33.8848 | -6.7627
3.2902 | -33.885 | -6.7765
3.2963 | -33.8852 | -6.7896
3.3024 | -33.8851 | -6.8013
3.3085 | -33.8854 | -6.815
3.3146 | -33.8853 | -6.8274
3.3207 | -33.8856 | -6.8408
3.3268 | -33.8858 | -6.8544
3.3329 | -33.8859 | -6.8657
3.339 | -33.886 | -6.8769
3.3451 | -33.8861 | -6.8896
3.3512 | -33.8863 | -6.9006
3.3573 | -33.8865 | -6.9133
3.3634 | -33.8866 | -6.9262
3.3696 | -33.8867 | -6.9394
3.3757 | -33.8867 | -6.9507
3.3818 | -33.8871 | -6.9637
3.3879 | -33.887 | -6.9756
3.394 | -33.8871 | -6.9875
3.4001 | -33.8872 | -6.9998
3.4062 | -33.8872 | -7.0115
3.4123 | -33.8873 | -7.0227
3.4184 | -33.8874 | -7.0352
3.4245 | -33.8876 | -7.0471
3.4306 | -33.8876 | -7.0583
3.4367 | -33.8877 | -7.0696
3.4428 | -33.8881 | -7.0803
3.4489 | -33.8881 | -7.0927
3.455 | -33.888 | -7.1046
3.4611 | -33.8881 | -7.1171
3.4672 | -33.8882 | -7.1298
3.4733 | -33.8883 | -7.1414
3.4794 | -33.8884 | -7.1533
3.4855 | -33.8886 | -7.1644
3.4916 | -33.8888 | -7.1756
3.4977 | -33.8888 | -7.188
3.5038 | -33.889 | -7.1994
3.5099 | -33.8891 | -7.2105
3.5161 | -33.8893 | -7.2215
3.5222 | -33.8895 | -7.2329
3.5283 | -33.8894 | -7.2445
3.5344 | -33.8895 | -7.2566
3.5405 | -33.8897 | -7.268
3.5466 | -33.8899 | -7.2792
3.5527 | -33.89 | -7.2913
3.5588 | -33.8901 | -7.3025
3.5649 | -33.8902 | -7.3141
3.571 | -33.8904 | -7.3257
3.5771 | -33.8905 | -7.3379
3.5832 | -33.8907 | -7.3489
3.5893 | -33.891 | -7.3602
3.5954 | -33.8911 | -7.3721
3.6015 | -33.8913 | -7.3839
3.6076 | -33.8916 | -7.395
3.6137 | -33.8917 | -7.4063
3.6198 | -33.8921 | -7.4168
3.6259 | -33.8922 | -7.429
3.632 | -33.8924 | -7.4409
3.6381 | -33.8927 | -7.4523
3.6442 | -33.8929 | -7.4641
3.6503 | -33.8932 | -7.4747
3.6565 | -33.8935 | -7.4876
3.6626 | -33.8937 | -7.4997
3.6687 | -33.894 | -7.5117
3.6748 | -33.8944 | -7.523
3.6809 | -33.8947 | -7.5349
3.687 | -33.895 | -7.5474
3.6931 | -33.8954 | -7.5587
3.6992 | -33.8956 | -7.5707
3.7053 | -33.8959 | -7.5822
3.7114 | -33.8961 | -7.5948
3.7175 | -33.8965 | -7.6061
3.7236 | -33.8968 | -7.6164
3.7297 | -33.8969 | -7.6285
3.7358 | -33.8972 | -7.6415
3.7419 | -33.8975 | -7.6543
3.748 | -33.8977 | -7.6664
3.7541 | -33.8978 | -7.6787
3.7602 | -33.8983 | -7.6908
3.7663 | -33.8985 | -7.7023
3.7724 | -33.8989 | -7.7147
3.7785 | -33.8992 | -7.7264
3.7846 | -33.8995 | -7.7384
3.7907 | -33.8999 | -7.7507
3.7969 | -33.9002 | -7.7635
3.803 | -33.9006 | -7.7749
3.8091 | -33.9009 | -7.7866
3.8152 | -33.9012 | -7.7991
3.8213 | -33.9015 | -7.8119
3.8274 | -33.9019 | -7.8238
3.8335 | -33.9023 | -7.8366
3.8396 | -33.9028 | -7.8494
3.8457 | -33.9032 | -7.861
3.8518 | -33.9037 | -7.8732
3.8579 | -33.9041 | -7.8851
3.864 | -33.9045 | -7.8972
3.8701 | -33.9049 | -7.9095
3.8762 | -33.9053 | -7.9219
3.8823 | -33.9056 | -7.9339
3.8884 | -33.906 | -7.9456
3.8945 | -33.9065 | -7.9582
3.9006 | -33.907 | -7.9711
3.9067 | -33.9074 | -7.9841
3.9128 | -33.9078 | -7.9967
3.9189 | -33.9082 | -8.0087
3.925 | -33.9085 | -8.0209
3.9311 | -33.9089 | -8.0323
3.9372 | -33.9094 | -8.0451
3.9434 | -33.9097 | -8.0569
3.9495 | -33.9102 | -8.0698
3.9556 | -33.9106 | -8.0823
3.9617 | -33.911 | -8.0944
3.9678 | -33.9115 | -8.1065
3.9739 | -33.9119 | -8.1189
3.98 | -33.9123 | -8.132
3.9861 | -33.9126 | -8.1438
3.9922 | -33.9129 | -8.1567
3.9983 | -33.9133 | -8.1682
4.0044 | -33.9137 | -8.1793
4.0105 | -33.9141 | -8.1915
4.0166 | -33.9146 | -8.2039
4.0227 | -33.9151 | -8.2156
4.0288 | -33.9153 | -8.2283
4.0349 | -33.9157 | -8.2396
4.041 | -33.9161 | -8.2518
4.0471 | -33.9164 | -8.2635
4.0532 | -33.9169 | -8.276
4.0593 | -33.9172 | -8.2888
4.0654 | -33.9176 | -8.3009
4.0715 | -33.9179 | -8.3136
4.0776 | -33.9183 | -8.3253
4.0838 | -33.9187 | -8.338
4.0899 | -33.9191 | -8.3505
4.096 | -33.9195 | -8.3623
4.1021 | -33.9198 | -8.3745
4.1082 | -33.9202 | -8.3861
4.1143 | -33.9206 | -8.3975
4.1204 | -33.9211 | -8.4096
4.1265 | -33.9215 | -8.4214
4.1326 | -33.9219 | -8.433
4.1387 | -33.9224 | -8.4456
4.1448 | -33.9227 | -8.4578
4.1509 | -33.923 | -8.4699
4.157 | -33.9233 | -8.4825
4.1631 | -33.9237 | -8.495
4.1692 | -33.9241 | -8.507
4.1753 | -33.9245 | -8.5192
4.1814 | -33.9248 | -8.5305
4.1875 | -33.9252 | -8.5432
4.1936 | -33.9255 | -8.5558
4.1997 | -33.9258 | -8.5673
4.2058 | -33.9261 | -8.5802
4.2119 | -33.9265 | -8.5913
4.218 | -33.9268 | -8.6029
4.2241 | -33.9272 | -8.6139
4.2303 | -33.9276 | -8.6265
4.2364 | -33.928 | -8.6391
4.2425 | -33.9283 | -8.651
4.2486 | -33.9286 | -8.6635
4.2547 | -33.9288 | -8.6742
4.2608 | -33.9291 | -8.6856
4.2669 | -33.9295 | -8.6974
4.273 | -33.9298 | -8.7095
4.2791 | -33.9301 | -8.7216
4.2852 | -33.9303 | -8.7333
4.2913 | -33.9305 | -8.744
4.2974 | -33.9308 | -8.7563
4.3035 | -33.9311 | -8.7685
4.3096 | -33.9313 | -8.7803
4.3157 | -33.9316 | -8.7917
4.3218 | -33.9319 | -8.8031
4.3279 | -33.9322 | -8.8149
4.334 | -33.9325 | -8.8261
4.3401 | -33.9327 | -8.8375
4.3462 | -33.9331 | -8.8496
4.3523 | -33.9335 | -8.8607
4.3584 | -33.9337 | -8.8708
4.3645 | -33.9339 | -8.8819
4.3707 | -33.9342 | -8.8927
4.3768 | -33.9345 | -8.9044
4.3829 | -33.9347 | -8.9162
4.389 | -33.935 | -8.9273
4.3951 | -33.9353 | -8.9383
4.4012 | -33.9355 | -8.9496
4.4073 | -33.9358 | -8.9611
4.4134 | -33.9361 | -8.9727
4.4195 | -33.9364 | -8.9838
4.4256 | -33.9367 | -8.9948
4.4317 | -33.937 | -9.0058
4.4378 | -33.9373 | -9.0168
4.4439 | -33.9375 | -9.0277
4.45 | -33.9377 | -9.0391
4.4561 | -33.9379 | -9.0504
4.4622 | -33.9381 | -9.0615
4.4683 | -33.9383 | -9.0722
4.4744 | -33.9386 | -9.0838
4.4805 | -33.9389 | -9.0944
4.4866 | -33.9392 | -9.1051
4.4927 | -33.9394 | -9.1167
4.4988 | -33.9397 | -9.1276
4.5049 | -33.9399 | -9.1382
4.511 | -33.9403 | -9.1487
4.5172 | -33.9406 | -9.16
4.5233 | -33.9408 | -9.172
4.5294 | -33.9411 | -9.1828
4.5355 | -33.9414 | -9.1937
4.5416 | -33.9417 | -9.2042
4.5477 | -33.942 | -9.215
4.5538 | -33.9422 | -9.2252
4.5599 | -33.9426 | -9.2362
4.566 | -33.9429 | -9.2473
4.5721 | -33.9431 | -9.258
4.5782 | -33.9435 | -9.2693
4.5843 | -33.9438 | -9.2803
4.5904 | -33.9441 | -9.2911
4.5965 | -33.9445 | -9.3014
4.6026 | -33.9448 | -9.3134
4.6087 | -33.945 | -9.3239
4.6148 | -33.9453 | -9.3359
4.6209 | -33.9458 | -9.3459
4.627 | -33.9461 | -9.3574
4.6331 | -33.9463 | -9.3678
4.6392 | -33.9466 | -9.3785
4.6453 | -33.9469 | -9.3893
4.6514 | -33.9473 | -9.4003
4.6576 | -33.9474 | -9.4122
4.6637 | -33.9477 | -9.4215
4.6698 | -33.9479 | -9.4316
4.6759 | -33.9481 | -9.4428
4.682 | -33.9484 | -9.4535
4.6881 | -33.9486 | -9.4638
4.6942 | -33.9489 | -9.4745
4.7003 | -33.9492 | -9.4854
4.7064 | -33.9495 | -9.4961
4.7125 | -33.9499 | -9.5071
4.7186 | -33.9501 | -9.517
4.7247 | -33.9505 | -9.5281
4.7308 | -33.9509 | -9.5389
4.7369 | -33.9511 | -9.5499
4.743 | -33.9514 | -9.5616
4.7491 | -33.9516 | -9.5727
4.7552 | -33.9518 | -9.5834
4.7613 | -33.9521 | -9.5945
4.7674 | -33.9525 | -9.6046
4.7735 | -33.9526 | -9.6161
4.7796 | -33.953 | -9.6273
4.7857 | -33.9531 | -9.6384
4.7918 | -33.9534 | -9.6486
4.7979 | -33.9538 | -9.6591
4.8041 | -33.954 | -9.6699
4.8102 | -33.9544 | -9.6804
4.8163 | -33.9547 | -9.6913
4.8224 | -33.9549 | -9.7015
4.8285 | -33.9553 | -9.712
4.8346 | -33.9556 | -9.723
4.8407 | -33.9559 | -9.7338
4.8468 | -33.9562 | -9.7449
4.8529 | -33.9565 | -9.7561
4.859 | -33.9567 | -9.7671
4.8651 | -33.957 | -9.7779
4.8712 | -33.9573 | -9.7887
4.8773 | -33.9576 | -9.7998
4.8834 | -33.9579 | -9.8109
4.8895 | -33.9583 | -9.8223
4.8956 | -33.9585 | -9.8336
4.9017 | -33.9587 | -9.8445
4.9078 | -33.959 | -9.8559
4.9139 | -33.9593 | -9.8659
4.92 | -33.9597 | -9.8764
4.9261 | -33.96 | -9.8867
4.9322 | -33.9605 | -9.8974
4.9383 | -33.9609 | -9.9079
4.9445 | -33.9612 | -9.9194
4.9506 | -33.9615 | -9.9297
4.9567 | -33.9618 | -9.9404
4.9628 | -33.9621 | -9.9513
4.9689 | -33.9624 | -9.9625
4.975 | -33.9627 | -9.9734
4.9811 | -33.963 | -9.9845
4.9872 | -33.9633 | -9.9957
4.9933 | -33.9638 | -10.0068
4.9994 | -33.9642 | -10.0181
5.0055 | -33.9646 | -10.0292
5.0116 | -33.9649 | -10.0408
5.0177 | -33.9654 | -10.0523
5.0238 | -33.9658 | -10.0641
5.0299 | -33.9663 | -10.0764
5.036 | -33.9666 | -10.0873
5.0421 | -33.9669 | -10.0985
5.0482 | -33.9672 | -10.1101
5.0543 | -33.9676 | -10.1208
5.0604 | -33.968 | -10.1324
5.0665 | -33.9684 | -10.1435
5.0726 | -33.9688 | -10.1552
5.0787 | -33.9692 | -10.1668
5.0848 | -33.9696 | -10.1783
5.091 | -33.9699 | -10.1895
5.0971 | -33.9703 | -10.201
5.1032 | -33.9707 | -10.2127
5.1093 | -33.9711 | -10.224
5.1154 | -33.9714 | -10.2354
5.1215 | -33.9717 | -10.2468
5.1276 | -33.9721 | -10.258
5.1337 | -33.9726 | -10.269
5.1398 | -33.9729 | -10.2802
5.1459 | -33.9734 | -10.292
5.152 | -33.9737 | -10.3043
5.1581 | -33.974 | -10.3163
5.1642 | -33.9744 | -10.3276
5.1703 | -33.9748 | -10.339
5.1764 | -33.9752 | -10.3503
5.1825 | -33.9755 | -10.3619
5.1886 | -33.9759 | -10.3731
5.1947 | -33.9762 | -10.3852
5.2008 | -33.9767 | -10.3964
5.2069 | -33.9771 | -10.4088
5.213 | -33.9775 | -10.4198
5.2191 | -33.978 | -10.4314
5.2252 | -33.9783 | -10.4431
5.2314 | -33.9786 | -10.4551
5.2375 | -33.9791 | -10.4664
5.2436 | -33.9794 | -10.4779
5.2497 | -33.9799 | -10.4891
5.2558 | -33.9803 | -10.5014
5.2619 | -33.9807 | -10.5136
5.268 | -33.9811 | -10.5254
5.2741 | -33.9815 | -10.5378
5.2802 | -33.9819 | -10.5499
5.2863 | -33.9824 | -10.5615
5.2924 | -33.9827 | -10.5739
5.2985 | -33.9832 | -10.5865
5.3046 | -33.9837 | -10.5989
5.3107 | -33.9841 | -10.611
5.3168 | -33.9846 | -10.6228
5.3229 | -33.985 | -10.634
5.329 | -33.9854 | -10.6454
5.3351 | -33.9859 | -10.6569
5.3412 | -33.9864 | -10.6684
5.3473 | -33.9868 | -10.6805
5.3534 | -33.9873 | -10.6927
5.3595 | -33.9877 | -10.7047
5.3656 | -33.9881 | -10.7169
5.3717 | -33.9885 | -10.7288
5.3779 | -33.9889 | -10.741
5.384 | -33.9893 | -10.753
5.3901 | -33.9897 | -10.765
5.3962 | -33.9901 | -10.7771
5.4023 | -33.9905 | -10.7891
5.4084 | -33.991 | -10.8008
5.4145 | -33.9914 | -10.813
5.4206 | -33.9917 | -10.8256
5.4267 | -33.9922 | -10.8381
5.4328 | -33.9926 | -10.8496
5.4389 | -33.993 | -10.8612
5.445 | -33.9934 | -10.8738
5.4511 | -33.9939 | -10.8854
5.4572 | -33.9943 | -10.8979
5.4633 | -33.9947 | -10.9096
5.4694 | -33.9951 | -10.9214
5.4755 | -33.9955 | -10.9335
5.4816 | -33.9959 | -10.9454
5.4877 | -33.9963 | -10.9573
5.4938 | -33.9967 | -10.969
5.4999 | -33.9971 | -10.981
5.506 | -33.9976 | -10.993
5.5121 | -33.998 | -11.0053
5.5183 | -33.9984 | -11.0167
5.5244 | -33.9987 | -11.0288
5.5305 | -33.9992 | -11.041
5.5366 | -33.9996 | -11.0527
5.5427 | -34 | -11.0652
5.5488 | -34.0003 | -11.0774
5.5549 | -34.0007 | -11.0894
5.561 | -34.0011 | -11.1019
5.5671 | -34.0016 | -11.1136
5.5732 | -34.0019 | -11.1258
5.5793 | -34.0024 | -11.1385
5.5854 | -34.0029 | -11.1507
5.5915 | -34.0032 | -11.1629
5.5976 | -34.0036 | -11.1752
5.6037 | -34.004 | -11.1871
5.6098 | -34.0043 | -11.1989
5.6159 | -34.0048 | -11.2111
5.622 | -34.0052 | -11.2228
5.6281 | -34.0056 | -11.2349
5.6342 | -34.006 | -11.2471
5.6403 | -34.0064 | -11.2587
5.6464 | -34.0069 | -11.2707
5.6525 | -34.0072 | -11.2825
5.6586 | -34.0077 | -11.2938
5.6648 | -34.008 | -11.3063
5.6709 | -34.0085 | -11.3187
5.677 | -34.0089 | -11.3313
5.6831 | -34.0092 | -11.3429
5.6892 | -34.0097 | -11.3549
5.6953 | -34.01 | -11.3668
5.7014 | -34.0105 | -11.3793
5.7075 | -34.0108 | -11.3916
5.7136 | -34.0112 | -11.4034
5.7197 | -34.0116 | -11.4152
5.7258 | -34.0119 | -11.427
5.7319 | -34.0124 | -11.439
5.738 | -34.0128 | -11.451
5.7441 | -34.0132 | -11.463
5.7502 | -34.0137 | -11.4746
5.7563 | -34.0141 | -11.4863
5.7624 | -34.0144 | -11.4989
5.7685 | -34.0148 | -11.5108
5.7746 | -34.0151 | -11.5226
5.7807 | -34.0155 | -11.5343
5.7868 | -34.0159 | -11.5462
5.7929 | -34.0162 | -11.558
5.799 | -34.0166 | -11.5699
5.8052 | -34.017 | -11.5819
5.8113 | -34.0173 | -11.5938
5.8174 | -34.0177 | -11.6059
5.8235 | -34.0181 | -11.6177
5.8296 | -34.0184 | -11.6301
5.8357 | -34.0187 | -11.642
5.8418 | -34.0192 | -11.6541
5.8479 | -34.0194 | -11.6663
5.854 | -34.0198 | -11.6785
5.8601 | -34.0202 | -11.6902
5.8662 | -34.0206 | -11.7025
5.8723 | -34.0209 | -11.7138
5.8784 | -34.0213 | -11.7257
5.8845 | -34.0217 | -11.7377
5.8906 | -34.022 | -11.7495
5.8967 | -34.0224 | -11.7611
5.9028 | -34.0229 | -11.773
5.9089 | -34.0232 | -11.7843
5.915 | -34.0236 | -11.7962
5.9211 | -34.024 | -11.8081
5.9272 | -34.0243 | -11.8199
5.9333 | -34.0247 | -11.8311
5.9394 | -34.0251 | -11.8428
5.9455 | -34.0255 | -11.8542
5.9517 | -34.026 | -11.8661
5.9578 | -34.0264 | -11.8782
5.9639 | -34.0268 | -11.8896
5.97 | -34.0272 | -11.9016
5.9761 | -34.0276 | -11.9131
5.9822 | -34.028 | -11.925
5.9883 | -34.0284 | -11.9366
5.9944 | -34.0288 | -11.9485
6.0005 | -34.0293 | -11.9605
6.0066 | -34.0296 | -11.9719
6.0127 | -34.03 | -11.983
6.0188 | -34.0304 | -11.9952
6.0249 | -34.0308 | -12.0069
6.031 | -34.0311 | -12.0184
6.0371 | -34.0316 | -12.0304
6.0432 | -34.032 | -12.0422
6.0493 | -34.0325 | -12.0537
6.0554 | -34.0329 | -12.0653
6.0615 | -34.0333 | -12.0775
6.0676 | -34.0337 | -12.0898
6.0737 | -34.0341 | -12.1018
6.0798 | -34.0344 | -12.1133
6.0859 | -34.0349 | -12.125
6.0921 | -34.0353 | -12.1365
6.0982 | -34.0357 | -12.1484
6.1043 | -34.0362 | -12.16
6.1104 | -34.0366 | -12.1722
6.1165 | -34.037 | -12.1832
6.1226 | -34.0375 | -12.1944
6.1287 | -34.0379 | -12.2061
6.1348 | -34.0384 | -12.2179
6.1409 | -34.0388 | -12.2295
6.147 | -34.0392 | -12.241
6.1531 | -34.0396 | -12.2526
6.1592 | -34.04 | -12.2643
6.1653 | -34.0404 | -12.2758
6.1714 | -34.0408 | -12.2875
6.1775 | -34.0413 | -12.2994
6.1836 | -34.0416 | -12.3109
6.1897 | -34.0421 | -12.3226
6.1958 | -34.0424 | -12.3345
6.2019 | -34.0429 | -12.3463
6.208 | -34.0433 | -12.3584
6.2141 | -34.0438 | -12.3702
6.2202 | -34.0442 | -12.3823
6.2263 | -34.0446 | -12.3938
6.2325 | -34.0451 | -12.405
6.2386 | -34.0455 | -12.4163
6.2447 | -34.0459 | -12.4282
6.2508 | -34.0463 | -12.4398
6.2569 | -34.0468 | -12.4517
6.263 | -34.0472 | -12.463
6.2691 | -34.0477 | -12.4745
6.2752 | -34.0481 | -12.4861
6.2813 | -34.0485 | -12.4977
6.2874 | -34.0489 | -12.5093
6.2935 | -34.0494 | -12.5207
6.2996 | -34.0498 | -12.5324
6.3057 | -34.0503 | -12.544
6.3118 | -34.0507 | -12.5555
6.3179 | -34.0511 | -12.5671
6.324 | -34.0515 | -12.579
6.3301 | -34.052 | -12.5904
6.3362 | -34.0525 | -12.602
6.3423 | -34.053 | -12.6137
6.3484 | -34.0534 | -12.6255
6.3545 | -34.0539 | -12.6366
6.3606 | -34.0543 | -12.6479
6.3667 | -34.0548 | -12.6593
6.3728 | -34.0553 | -12.6709
6.379 | -34.0557 | -12.6824
6.3851 | -34.0561 | -12.694
6.3912 | -34.0566 | -12.7054
6.3973 | -34.0571 | -12.7167
6.4034 | -34.0575 | -12.7282
6.4095 | -34.0579 | -12.74
6.4156 | -34.0583 | -12.7512
6.4217 | -34.0587 | -12.7628
6.4278 | -34.0592 | -12.7742
6.4339 | -34.0597 | -12.7857
6.44 | -34.0601 | -12.7969
6.4461 | -34.0605 | -12.8089
6.4522 | -34.061 | -12.8199
6.4583 | -34.0614 | -12.8315
6.4644 | -34.0618 | -12.8434
6.4705 | -34.0623 | -12.855
6.4766 | -34.0627 | -12.8669
6.4827 | -34.0632 | -12.8788
6.4888 | -34.0636 | -12.8902
6.4949 | -34.064 | -12.9013
6.501 | -34.0644 | -12.9131
6.5071 | -34.0649 | -12.9247
6.5132 | -34.0654 | -12.9361
6.5194 | -34.0658 | -12.9476
6.5255 | -34.0663 | -12.9587
6.5316 | -34.0667 | -12.9704
6.5377 | -34.0672 | -12.9822
6.5438 | -34.0676 | -12.9936
6.5499 | -34.0681 | -13.0049
6.556 | -34.0685 | -13.0162
6.5621 | -34.0689 | -13.0279
6.5682 | -34.0694 | -13.0393
6.5743 | -34.0698 | -13.0515
6.5804 | -34.0702 | -13.063
6.5865 | -34.0706 | -13.0747
6.5926 | -34.0711 | -13.0862
6.5987 | -34.0716 | -13.0973
6.6048 | -34.072 | -13.1094
6.6109 | -34.0725 | -13.1209
6.617 | -34.0729 | -13.1332
6.6231 | -34.0732 | -13.1449
6.6292 | -34.0738 | -13.1563
6.6353 | -34.0742 | -13.1687
6.6414 | -34.0747 | -13.1805
6.6475 | -34.0752 | -13.1922
6.6536 | -34.0756 | -13.2034
6.6597 | -34.0761 | -13.2147
6.6659 | -34.0766 | -13.2261
6.672 | -34.077 | -13.2377
6.6781 | -34.0775 | -13.249
6.6842 | -34.0779 | -13.2607
6.6903 | -34.0784 | -13.2721
6.6964 | -34.0789 | -13.2837
6.7025 | -34.0793 | -13.2952
6.7086 | -34.0797 | -13.3065
6.7147 | -34.0802 | -13.3182
6.7208 | -34.0806 | -13.3302
6.7269 | -34.081 | -13.3413
6.733 | -34.0814 | -13.3521
6.7391 | -34.0819 | -13.3639
6.7452 | -34.0824 | -13.3748
6.7513 | -34.0829 | -13.386
6.7574 | -34.0834 | -13.3978
6.7635 | -34.0838 | -13.4088
6.7696 | -34.0844 | -13.4204
6.7757 | -34.0848 | -13.4321
6.7818 | -34.0853 | -13.4437
6.7879 | -34.0858 | -13.4552
6.794 | -34.0862 | -13.4665
6.8001 | -34.0866 | -13.4778
6.8063 | -34.0871 | -13.4895
6.8124 | -34.0875 | -13.5011
6.8185 | -34.088 | -13.5125
6.8246 | -34.0884 | -13.524
6.8307 | -34.0889 | -13.5353
6.8368 | -34.0893 | -13.5471
6.8429 | -34.0898 | -13.5585
6.849 | -34.0902 | -13.57
6.8551 | -34.0907 | -13.5815
6.8612 | -34.0912 | -13.5928
6.8673 | -34.0917 | -13.6041
6.8734 | -34.0921 | -13.6162
6.8795 | -34.0926 | -13.6279
6.8856 | -34.0931 | -13.6389
6.8917 | -34.0935 | -13.6499
6.8978 | -34.0939 | -13.6613
6.9039 | -34.0944 | -13.6725
6.91 | -34.0949 | -13.6841
6.9161 | -34.0953 | -13.6953
6.9222 | -34.0957 | -13.7068
6.9283 | -34.0962 | -13.7179
6.9344 | -34.0966 | -13.7291
6.9405 | -34.0971 | -13.7403
6.9466 | -34.0976 | -13.752
6.9528 | -34.098 | -13.7635
6.9589 | -34.0985 | -13.7747
6.965 | -34.0989 | -13.7862
6.9711 | -34.0994 | -13.7976
6.9772 | -34.0999 | -13.8088
6.9833 | -34.1003 | -13.8203
6.9894 | -34.1007 | -13.8321
6.9955 | -34.1012 | -13.8435
7.0016 | -34.1016 | -13.8552
7.0077 | -34.1021 | -13.8669
7.0138 | -34.1026 | -13.8785
7.0199 | -34.103 | -13.8897
7.026 | -34.1034 | -13.901
7.0321 | -34.1038 | -13.9124
7.0382 | -34.1042 | -13.9235
7.0443 | -34.1048 | -13.9352
7.0504 | -34.1052 | -13.9469
7.0565 | -34.1056 | -13.9585
7.0626 | -34.1061 | -13.9697
7.0687 | -34.1065 | -13.9806
7.0748 | -34.107 | -13.9918
7.0809 | -34.1074 | -14.0033
7.087 | -34.1078 | -14.0148
7.0932 | -34.1082 | -14.0261
7.0993 | -34.1087 | -14.0375
7.1054 | -34.1091 | -14.0486
7.1115 | -34.1095 | -14.0598
7.1176 | -34.1099 | -14.071
7.1237 | -34.1103 | -14.0824
7.1298 | -34.1107 | -14.0933
7.1359 | -34.1111 | -14.1044
7.142 | -34.1115 | -14.1157
7.1481 | -34.1119 | -14.1268
7.1542 | -34.1124 | -14.138
7.1603 | -34.1128 | -14.1492
7.1664 | -34.1133 | -14.1605
7.1725 | -34.1137 | -14.1716
7.1786 | -34.1142 | -14.183
7.1847 | -34.1146 | -14.1941
7.1908 | -34.115 | -14.2055
7.1969 | -34.1154 | -14.2168
7.203 | -34.1159 | -14.2281
7.2091 | -34.1163 | -14.2391
7.2152 | -34.1167 | -14.2504
7.2213 | -34.1172 | -14.2614
7.2274 | -34.1176 | -14.2727
7.2335 | -34.118 | -14.2839
7.2397 | -34.1185 | -14.2951
7.2458 | -34.1189 | -14.3059
7.2519 | -34.1193 | -14.3171
7.258 | -34.1197 | -14.3285
7.2641 | -34.1202 | -14.3396
7.2702 | -34.1206 | -14.3511
7.2763 | -34.1211 | -14.3621
7.2824 | -34.1216 | -14.3732
7.2885 | -34.122 | -14.3844
7.2946 | -34.1224 | -14.3957
7.3007 | -34.1228 | -14.407
7.3068 | -34.1232 | -14.4184
7.3129 | -34.1236 | -14.4294
7.319 | -34.1241 | -14.4407
7.3251 | -34.1245 | -14.4515
7.3312 | -34.125 | -14.4624
7.3373 | -34.1255 | -14.4733
7.3434 | -34.126 | -14.4842
7.3495 | -34.1264 | -14.495
7.3556 | -34.1268 | -14.5065
7.3617 | -34.1272 | -14.518
7.3678 | -34.1277 | -14.5293
7.3739 | -34.1281 | -14.5401
7.3801 | -34.1285 | -14.5511
7.3862 | -34.1289 | -14.5622
7.3923 | -34.1294 | -14.5733
7.3984 | -34.1298 | -14.5843
7.4045 | -34.1302 | -14.5953
7.4106 | -34.1307 | -14.6063
7.4167 | -34.1311 | -14.6176
7.4228 | -34.1316 | -14.6286
7.4289 | -34.132 | -14.6395
7.435 | -34.1325 | -14.6503
7.4411 | -34.1329 | -14.6611
7.4472 | -34.1333 | -14.6723
7.4533 | -34.1338 | -14.6834
7.4594 | -34.1342 | -14.6947
7.4655 | -34.1347 | -14.7057
7.4716 | -34.1351 | -14.717
7.4777 | -34.1355 | -14.7276
7.4838 | -34.136 | -14.7388
7.4899 | -34.1365 | -14.7498
7.496 | -34.1369 | -14.7609
7.5021 | -34.1374 | -14.7718
7.5082 | -34.1378 | -14.783
7.5143 | -34.1382 | -14.7939
7.5204 | -34.1386 | -14.8047
7.5266 | -34.139 | -14.8157
7.5327 | -34.1395 | -14.8267
7.5388 | -34.1399 | -14.8377
7.5449 | -34.1403 | -14.8489
7.551 | -34.1407 | -14.8601
7.5571 | -34.1412 | -14.871
7.5632 | -34.1416 | -14.882
7.5693 | -34.1421 | -14.8931
7.5754 | -34.1425 | -14.9041
7.5815 | -34.143 | -14.9148
7.5876 | -34.1434 | -14.9258
7.5937 | -34.1438 | -14.9367
7.5998 | -34.1442 | -14.9478
7.6059 | -34.1446 | -14.9588
7.612 | -34.1451 | -14.9698
7.6181 | -34.1455 | -14.9807
7.6242 | -34.146 | -14.9918
7.6303 | -34.1465 | -15.0026
7.6364 | -34.1469 | -15.0138
7.6425 | -34.1473 | -15.0251
7.6486 | -34.1477 | -15.0361
7.6547 | -34.1482 | -15.0472
7.6608 | -34.1487 | -15.0581
7.667 | -34.1491 | -15.0686
7.6731 | -34.1495 | -15.0798
7.6792 | -34.1499 | -15.0909
7.6853 | -34.1503 | -15.1018
7.6914 | -34.1506 | -15.1129
7.6975 | -34.1511 | -15.1241
7.7036 | -34.1515 | -15.1351
7.7097 | -34.152 | -15.1461
7.7158 | -34.1524 | -15.157
7.7219 | -34.1528 | -15.1679
7.728 | -34.1532 | -15.179
7.7341 | -34.1536 | -15.19
7.7402 | -34.1541 | -15.2009
7.7463 | -34.1545 | -15.2119
7.7524 | -34.1549 | -15.2232
7.7585 | -34.1553 | -15.2343
7.7646 | -34.1557 | -15.245
7.7707 | -34.1561 | -15.2558
7.7768 | -34.1564 | -15.2667
7.7829 | -34.1568 | -15.2778
7.789 | -34.1572 | -15.2886
7.7951 | -34.1576 | -15.2999
7.8012 | -34.1581 | -15.3112
7.8073 | -34.1585 | -15.322
7.8135 | -34.1589 | -15.3331
7.8196 | -34.1593 | -15.344
7.8257 | -34.1598 | -15.3554
7.8318 | -34.1602 | -15.3667
7.8379 | -34.1606 | -15.3773
7.844 | -34.161 | -15.3879
7.8501 | -34.1613 | -15.3991
7.8562 | -34.1617 | -15.4107
7.8623 | -34.1622 | -15.4217
7.8684 | -34.1625 | -15.4324
7.8745 | -34.163 | -15.4432
7.8806 | -34.1634 | -15.4538
7.8867 | -34.1638 | -15.4646
7.8928 | -34.1643 | -15.4754
7.8989 | -34.1647 | -15.486
7.905 | -34.1652 | -15.4973
7.9111 | -34.1656 | -15.5083
7.9172 | -34.166 | -15.519
7.9233 | -34.1663 | -15.53
7.9294 | -34.1668 | -15.5411
7.9355 | -34.1672 | -15.5516
7.9416 | -34.1676 | -15.5625
7.9477 | -34.168 | -15.573
7.9539 | -34.1684 | -15.5836
7.96 | -34.1688 | -15.5948
7.9661 | -34.1692 | -15.6058
7.9722 | -34.1697 | -15.6168
7.9783 | -34.1701 | -15.6277
7.9844 | -34.1706 | -15.639
7.9905 | -34.1709 | -15.6498
7.9966 | -34.1713 | -15.6608
8.0027 | -34.1717 | -15.672
8.0088 | -34.1721 | -15.6826
8.0149 | -34.1725 | -15.6932
8.021 | -34.173 | -15.7042
8.0271 | -34.1734 | -15.7155
8.0332 | -34.1738 | -15.7264
8.0393 | -34.1742 | -15.7375
8.0454 | -34.1746 | -15.7486
8.0515 | -34.175 | -15.7595
8.0576 | -34.1754 | -15.7707
8.0637 | -34.1758 | -15.7811
8.0698 | -34.1762 | -15.7917
8.0759 | -34.1766 | -15.8023
8.082 | -34.1771 | -15.8134
8.0881 | -34.1774 | -15.8242
8.0942 | -34.1779 | -15.8355
8.1004 | -34.1783 | -15.8467
8.1065 | -34.1787 | -15.8576
8.1126 | -34.1791 | -15.8684
8.1187 | -34.1796 | -15.8795
8.1248 | -34.18 | -15.8905
8.1309 | -34.1804 | -15.9014
8.137 | -34.1808 | -15.9122
8.1431 | -34.1812 | -15.9233
8.1492 | -34.1816 | -15.9343
8.1553 | -34.1821 | -15.9452
8.1614 | -34.1825 | -15.9564
8.1675 | -34.1828 | -15.9671
8.1736 | -34.1833 | -15.9783
8.1797 | -34.1836 | -15.9894
8.1858 | -34.184 | -16.0003
8.1919 | -34.1844 | -16.0112
8.198 | -34.1848 | -16.0217
8.2041 | -34.1853 | -16.033
8.2102 | -34.1857 | -16.0439
8.2163 | -34.1862 | -16.0548
8.2224 | -34.1866 | -16.0659
8.2285 | -34.187 | -16.0766
8.2346 | -34.1874 | -16.0874
8.2408 | -34.1878 | -16.0984
8.2469 | -34.1882 | -16.1092
8.253 | -34.1886 | -16.12
8.2591 | -34.1891 | -16.1307
8.2652 | -34.1895 | -16.1415
8.2713 | -34.19 | -16.1524
8.2774 | -34.1904 | -16.1634
8.2835 | -34.1908 | -16.1747
8.2896 | -34.1913 | -16.1851
8.2957 | -34.1917 | -16.1961
8.3018 | -34.1921 | -16.2069
8.3079 | -34.1926 | -16.2178
8.314 | -34.193 | -16.2286
8.3201 | -34.1934 | -16.2394
8.3262 | -34.1938 | -16.2502
8.3323 | -34.1942 | -16.261
8.3384 | -34.1947 | -16.2723
8.3445 | -34.1951 | -16.2828
8.3506 | -34.1955 | -16.2939
8.3567 | -34.1959 | -16.305
8.3628 | -34.1963 | -16.3158
8.3689 | -34.1968 | -16.3267
8.375 | -34.1972 | -16.3376
8.3812 | -34.1976 | -16.3485
8.3873 | -34.198 | -16.3595
8.3934 | -34.1984 | -16.3703
8.3995 | -34.1988 | -16.3812
8.4056 | -34.1993 | -16.3922
8.4117 | -34.1997 | -16.4033
8.4178 | -34.2001 | -16.4138
8.4239 | -34.2005 | -16.4245
8.43 | -34.201 | -16.4354
8.4361 | -34.2014 | -16.4466
8.4422 | -34.2018 | -16.4574
8.4483 | -34.2022 | -16.4683
8.4544 | -34.2027 | -16.4794
8.4605 | -34.2031 | -16.4902
8.4666 | -34.2034 | -16.5013
8.4727 | -34.2038 | -16.5121
8.4788 | -34.2043 | -16.5228
8.4849 | -34.2047 | -16.5339
8.491 | -34.2051 | -16.5447
8.4971 | -34.2055 | -16.5556
8.5032 | -34.206 | -16.5665
8.5093 | -34.2064 | -16.5773
8.5154 | -34.2068 | -16.588
8.5215 | -34.2072 | -16.599
8.5277 | -34.2076 | -16.6099
8.5338 | -34.208 | -16.6206
8.5399 | -34.2085 | -16.6317
8.546 | -34.2089 | -16.6423
8.5521 | -34.2093 | -16.6533
8.5582 | -34.2098 | -16.6643
8.5643 | -34.2102 | -16.675
8.5704 | -34.2106 | -16.6855
8.5765 | -34.211 | -16.6964
8.5826 | -34.2115 | -16.7072
8.5887 | -34.2119 | -16.7179
8.5948 | -34.2123 | -16.7284
8.6009 | -34.2128 | -16.7392
8.607 | -34.2132 | -16.7503
8.6131 | -34.2136 | -16.7614
8.6192 | -34.214 | -16.7719
8.6253 | -34.2145 | -16.7828
8.6314 | -34.2149 | -16.7936
8.6375 | -34.2153 | -16.8047
8.6436 | -34.2158 | -16.8156
8.6497 | -34.2163 | -16.8265
8.6558 | -34.2166 | -16.8371
8.6619 | -34.2171 | -16.8481
8.6681 | -34.2174 | -16.8589
8.6742 | -34.2179 | -16.87
8.6803 | -34.2184 | -16.881
8.6864 | -34.2188 | -16.8921
8.6925 | -34.2193 | -16.9029
8.6986 | -34.2197 | -16.914
8.7047 | -34.2201 | -16.925
8.7108 | -34.2205 | -16.9359
8.7169 | -34.221 | -16.9466
8.723 | -34.2215 | -16.9574
8.7291 | -34.2219 | -16.9681
8.7352 | -34.2223 | -16.9792
8.7413 | -34.2227 | -16.99
8.7474 | -34.2232 | -17.0006
8.7535 | -34.2237 | -17.0115
8.7596 | -34.2241 | -17.0229
8.7657 | -34.2246 | -17.0337
8.7718 | -34.225 | -17.0443
8.7779 | -34.2254 | -17.0557
8.784 | -34.2258 | -17.0664
8.7901 | -34.2262 | -17.0776
8.7962 | -34.2267 | -17.0885
8.8023 | -34.2272 | -17.0991
8.8084 | -34.2276 | -17.1101
8.8146 | -34.228 | -17.1208
8.8207 | -34.2285 | -17.1317
8.8268 | -34.229 | -17.1424
8.8329 | -34.2294 | -17.1533
8.839 | -34.2299 | -17.1642
8.8451 | -34.2303 | -17.1749
8.8512 | -34.2308 | -17.1858
8.8573 | -34.2312 | -17.1965
8.8634 | -34.2317 | -17.2072
8.8695 | -34.2321 | -17.2182
8.8756 | -34.2326 | -17.2289
8.8817 | -34.233 | -17.2399
8.8878 | -34.2334 | -17.251
8.8939 | -34.2338 | -17.2616
8.9 | -34.2343 | -17.2721
8.9061 | -34.2348 | -17.283
8.9122 | -34.2352 | -17.2937
8.9183 | -34.2356 | -17.3044
8.9244 | -34.2361 | -17.3151
8.9305 | -34.2366 | -17.3259
8.9366 | -34.237 | -17.3366
8.9427 | -34.2375 | -17.3475
8.9488 | -34.238 | -17.3583
8.955 | -34.2384 | -17.3691
8.9611 | -34.2389 | -17.3801
8.9672 | -34.2393 | -17.3912
8.9733 | -34.2398 | -17.4018
8.9794 | -34.2403 | -17.4128
8.9855 | -34.2407 | -17.4235
8.9916 | -34.2412 | -17.4342
8.9977 | -34.2416 | -17.445
9.0038 | -34.2421 | -17.4559
9.0099 | -34.2426 | -17.4668
9.016 | -34.2431 | -17.4776
9.0221 | -34.2435 | -17.4887
9.0282 | -34.244 | -17.4996
9.0343 | -34.2445 | -17.5106
9.0404 | -34.245 | -17.5213
9.0465 | -34.2455 | -17.5324
9.0526 | -34.246 | -17.5432
9.0587 | -34.2464 | -17.554
9.0648 | -34.2469 | -17.5647
9.0709 | -34.2474 | -17.5753
9.077 | -34.2479 | -17.5866
9.0831 | -34.2483 | -17.5974
9.0892 | -34.2487 | -17.6082
9.0953 | -34.2492 | -17.6191
9.1015 | -34.2496 | -17.63
9.1076 | -34.2502 | -17.6408
9.1137 | -34.2506 | -17.6515
9.1198 | -34.2511 | -17.662
9.1259 | -34.2516 | -17.6729
9.132 | -34.2521 | -17.684
9.1381 | -34.2526 | -17.6947
9.1442 | -34.2532 | -17.705
9.1503 | -34.2536 | -17.716
9.1564 | -34.2541 | -17.7268
9.1625 | -34.2545 | -17.7377
9.1686 | -34.255 | -17.7485
9.1747 | -34.2555 | -17.7594
9.1808 | -34.256 | -17.7703
9.1869 | -34.2564 | -17.7814
9.193 | -34.2569 | -17.7923
9.1991 | -34.2573 | -17.8029
9.2052 | -34.2579 | -17.8136
9.2113 | -34.2584 | -17.8246
9.2174 | -34.2589 | -17.836
9.2235 | -34.2594 | -17.8465
9.2296 | -34.2599 | -17.8574
9.2357 | -34.2603 | -17.8683
9.2419 | -34.2608 | -17.8791
9.248 | -34.2614 | -17.8901
9.2541 | -34.2618 | -17.901
9.2602 | -34.2623 | -17.912
9.2663 | -34.2628 | -17.9226
9.2724 | -34.2633 | -17.9335
9.2785 | -34.2637 | -17.9443
9.2846 | -34.2643 | -17.9552
9.2907 | -34.2647 | -17.9659
9.2968 | -34.2652 | -17.9768
9.3029 | -34.2657 | -17.9878
9.309 | -34.2662 | -17.9986
9.3151 | -34.2666 | -18.0092
9.3212 | -34.2672 | -18.0202
9.3273 | -34.2677 | -18.0309
9.3334 | -34.2682 | -18.0415
9.3395 | -34.2688 | -18.0522
9.3456 | -34.2693 | -18.0628
9.3517 | -34.2698 | -18.0733
9.3578 | -34.2703 | -18.0839
9.3639 | -34.2708 | -18.0948
9.37 | -34.2713 | -18.1057
9.3761 | -34.2718 | -18.1164
9.3822 | -34.2723 | -18.1271
9.3884 | -34.2728 | -18.1378
9.3945 | -34.2733 | -18.1487
9.4006 | -34.2738 | -18.1595
9.4067 | -34.2743 | -18.1704
9.4128 | -34.2748 | -18.1811
9.4189 | -34.2753 | -18.1921
9.425 | -34.2759 | -18.2027
9.4311 | -34.2763 | -18.2134
9.4372 | -34.2768 | -18.2242
9.4433 | -34.2773 | -18.2351
9.4494 | -34.2778 | -18.2458
9.4555 | -34.2783 | -18.2565
9.4616 | -34.2789 | -18.2672
9.4677 | -34.2794 | -18.2778
9.4738 | -34.2799 | -18.2889
9.4799 | -34.2804 | -18.2996
9.486 | -34.2808 | -18.3107
9.4921 | -34.2814 | -18.3218
9.4982 | -34.2819 | -18.3322
9.5043 | -34.2824 | -18.3427
9.5104 | -34.2829 | -18.3533
9.5165 | -34.2834 | -18.3641
9.5226 | -34.2839 | -18.375
9.5288 | -34.2845 | -18.3857
9.5349 | -34.285 | -18.3965
9.541 | -34.2855 | -18.4072
9.5471 | -34.286 | -18.418
9.5532 | -34.2865 | -18.4287
9.5593 | -34.2871 | -18.4395
9.5654 | -34.2876 | -18.4503
9.5715 | -34.2881 | -18.4613
9.5776 | -34.2886 | -18.4723
9.5837 | -34.2891 | -18.4827
9.5898 | -34.2896 | -18.4937
9.5959 | -34.29 | -18.5043
9.602 | -34.2905 | -18.5151
9.6081 | -34.2911 | -18.526
9.6142 | -34.2916 | -18.5366
9.6203 | -34.2921 | -18.5474
9.6264 | -34.2925 | -18.5578
9.6325 | -34.293 | -18.5687
9.6386 | -34.2935 | -18.5794
9.6447 | -34.294 | -18.5901
9.6508 | -34.2946 | -18.6009
9.6569 | -34.2951 | -18.6116
9.663 | -34.2956 | -18.6224
9.6691 | -34.2962 | -18.6331
9.6753 | -34.2967 | -18.6438
9.6814 | -34.2972 | -18.6543
9.6875 | -34.2978 | -18.6649
9.6936 | -34.2983 | -18.6755
9.6997 | -34.2988 | -18.6864
9.7058 | -34.2994 | -18.6971
9.7119 | -34.2999 | -18.7078
9.718 | -34.3004 | -18.7188
9.7241 | -34.3009 | -18.7295
9.7302 | -34.3014 | -18.7402
9.7363 | -34.3019 | -18.7508
9.7424 | -34.3025 | -18.7613
9.7485 | -34.3029 | -18.7719
9.7546 | -34.3035 | -18.7823
9.7607 | -34.304 | -18.7932
9.7668 | -34.3045 | -18.8043
9.7729 | -34.305 | -18.8148
9.779 | -34.3057 | -18.8255
9.7851 | -34.3061 | -18.8364
9.7912 | -34.3066 | -18.8476
9.7973 | -34.3071 | -18.8584
9.8034 | -34.3076 | -18.8693
9.8095 | -34.3081 | -18.8799
9.8157 | -34.3087 | -18.8906
9.8218 | -34.3092 | -18.901
9.8279 | -34.3097 | -18.9117
9.834 | -34.3103 | -18.9223
9.8401 | -34.3108 | -18.9329
9.8462 | -34.3113 | -18.9433
9.8523 | -34.3119 | -18.954
9.8584 | -34.3124 | -18.9649
9.8645 | -34.3129 | -18.9757
9.8706 | -34.3134 | -18.9862
9.8767 | -34.3139 | -18.9966
9.8828 | -34.3144 | -19.0069
9.8889 | -34.315 | -19.0177
9.895 | -34.3155 | -19.0281
9.9011 | -34.316 | -19.0384
9.9072 | -34.3165 | -19.0488
9.9133 | -34.317 | -19.0592
9.9194 | -34.3176 | -19.0697
9.9255 | -34.3181 | -19.0805
9.9316 | -34.3186 | -19.0909
9.9377 | -34.3192 | -19.1014
9.9438 | -34.3197 | -19.112
9.9499 | -34.3202 | -19.1228
9.956 | -34.3207 | -19.1332
9.9622 | -34.3213 | -19.1438
9.9683 | -34.3218 | -19.1544
9.9744 | -34.3223 | -19.165
9.9805 | -34.3229 | -19.1757
9.9866 | -34.3234 | -19.186
9.9927 | -34.3239 | -19.1967
9.9988 | -34.3245 | -19.2072
10.0049 | -34.325 | -19.2177
10.011 | -34.3255 | -19.2282
10.0171 | -34.326 | -19.2388
10.0232 | -34.3266 | -19.2491
10.0293 | -34.3272 | -19.2596
10.0354 | -34.3278 | -19.2699
10.0415 | -34.3283 | -19.2804
10.0476 | -34.3288 | -19.2913
10.0537 | -34.3293 | -19.3021
10.0598 | -34.3298 | -19.3128
10.0659 | -34.3303 | -19.3234
10.072 | -34.3309 | -19.3339
10.0781 | -34.3314 | -19.3444
10.0842 | -34.332 | -19.3551
10.0903 | -34.3325 | -19.3659
10.0964 | -34.3331 | -19.3761
10.1026 | -34.3337 | -19.3866
10.1087 | -34.3342 | -19.3972
10.1148 | -34.3347 | -19.4079
10.1209 | -34.3353 | -19.4183
10.127 | -34.3358 | -19.4286
10.1331 | -34.3364 | -19.439
10.1392 | -34.3369 | -19.4495
10.1453 | -34.3375 | -19.46
10.1514 | -34.338 | -19.4704
10.1575 | -34.3385 | -19.4808
10.1636 | -34.3391 | -19.4906
10.1697 | -34.3396 | -19.5012
10.1758 | -34.3402 | -19.5115
10.1819 | -34.3407 | -19.5221
10.188 | -34.3412 | -19.5324
10.1941 | -34.3417 | -19.5431
10.2002 | -34.3422 | -19.5531
10.2063 | -34.3427 | -19.5637
10.2124 | -34.3433 | -19.5742
10.2185 | -34.3438 | -19.5846
10.2246 | -34.3443 | -19.595
10.2307 | -34.3449 | -19.6057
10.2368 | -34.3454 | -19.6153
10.2429 | -34.346 | -19.6255
10.2491 | -34.3465 | -19.6357
10.2552 | -34.347 | -19.6465
10.2613 | -34.3476 | -19.6572
10.2674 | -34.3481 | -19.6678
10.2735 | -34.3486 | -19.6784
10.2796 | -34.3492 | -19.689
10.2857 | -34.3497 | -19.6995
10.2918 | -34.3502 | -19.71
10.2979 | -34.3508 | -19.7206
10.304 | -34.3514 | -19.7312
10.3101 | -34.3519 | -19.7416
10.3162 | -34.3525 | -19.752
10.3223 | -34.353 | -19.7624
10.3284 | -34.3535 | -19.7728
10.3345 | -34.354 | -19.783
10.3406 | -34.3546 | -19.7934
10.3467 | -34.3551 | -19.8038
10.3528 | -34.3556 | -19.814
10.3589 | -34.3561 | -19.8246
10.365 | -34.3567 | -19.8353
10.3711 | -34.3572 | -19.846
10.3772 | -34.3577 | -19.8564
10.3833 | -34.3582 | -19.8668
10.3895 | -34.3587 | -19.8771
10.3956 | -34.3593 | -19.8873
10.4017 | -34.3598 | -19.8976
10.4078 | -34.3603 | -19.9081
10.4139 | -34.3609 | -19.9185
10.42 | -34.3615 | -19.929
10.4261 | -34.362 | -19.9393
10.4322 | -34.3625 | -19.9495
10.4383 | -34.3631 | -19.9598
10.4444 | -34.3636 | -19.9698
10.4505 | -34.3642 | -19.9802
10.4566 | -34.3647 | -19.9905
10.4627 | -34.3652 | -20.0009
10.4688 | -34.3658 | -20.0108
10.4749 | -34.3663 | -20.0211
10.481 | -34.3669 | -20.0315
10.4871 | -34.3674 | -20.0419
10.4932 | -34.368 | -20.0522
10.4993 | -34.3686 | -20.0625
10.5054 | -34.3691 | -20.0728
10.5115 | -34.3696 | -20.0832
10.5176 | -34.3702 | -20.0932
10.5237 | -34.3707 | -20.1034
10.5298 | -34.3712 | -20.1135
10.536 | -34.3718 | -20.124
10.5421 | -34.3724 | -20.1339
10.5482 | -34.3729 | -20.1445
10.5543 | -34.3735 | -20.1548
10.5604 | -34.3741 | -20.1648
10.5665 | -34.3746 | -20.175
10.5726 | -34.3751 | -20.1851
10.5787 | -34.3757 | -20.1955
10.5848 | -34.3763 | -20.2053
10.5909 | -34.3768 | -20.2157
10.597 | -34.3773 | -20.2259
10.6031 | -34.3779 | -20.2364
10.6092 | -34.3784 | -20.2469
10.6153 | -34.3789 | -20.2574
10.6214 | -34.3795 | -20.2676
10.6275 | -34.3801 | -20.2777
10.6336 | -34.3806 | -20.2882
10.6397 | -34.3811 | -20.2983
10.6458 | -34.3816 | -20.3084
10.6519 | -34.3822 | -20.3185
10.658 | -34.3827 | -20.3284
10.6641 | -34.3832 | -20.3387
10.6702 | -34.3838 | -20.3493
10.6764 | -34.3844 | -20.3594
10.6825 | -34.3849 | -20.3698
10.6886 | -34.3855 | -20.3801
10.6947 | -34.386 | -20.3901
10.7008 | -34.3866 | -20.4003
10.7069 | -34.3872 | -20.4106
10.713 | -34.3877 | -20.4207
10.7191 | -34.3883 | -20.431
10.7252 | -34.3889 | -20.441
10.7313 | -34.3894 | -20.4512
10.7374 | -34.39 | -20.4611
10.7435 | -34.3905 | -20.4716
10.7496 | -34.3911 | -20.4815
10.7557 | -34.3916 | -20.4918
10.7618 | -34.3922 | -20.5019
10.7679 | -34.3927 | -20.512
10.774 | -34.3932 | -20.522
10.7801 | -34.3938 | -20.5325
10.7862 | -34.3943 | -20.5426
10.7923 | -34.3949 | -20.5528
10.7984 | -34.3954 | -20.5628
10.8045 | -34.396 | -20.5731
10.8106 | -34.3965 | -20.5833
10.8168 | -34.3971 | -20.5934
10.8229 | -34.3976 | -20.6036
10.829 | -34.3982 | -20.6139
10.8351 | -34.3987 | -20.6239
10.8412 | -34.3992 | -20.6339
10.8473 | -34.3998 | -20.6438
10.8534 | -34.4003 | -20.6542
10.8595 | -34.4009 | -20.6641
10.8656 | -34.4015 | -20.674
10.8717 | -34.402 | -20.6838
10.8778 | -34.4026 | -20.6938
10.8839 | -34.4031 | -20.7038
10.89 | -34.4036 | -20.7138
10.8961 | -34.4042 | -20.724
10.9022 | -34.4048 | -20.7341
10.9083 | -34.4054 | -20.7445
10.9144 | -34.4059 | -20.7544
10.9205 | -34.4065 | -20.7645
10.9266 | -34.4071 | -20.7749
10.9327 | -34.4077 | -20.7849
10.9388 | -34.4082 | -20.795
10.9449 | -34.4088 | -20.8051
10.951 | -34.4093 | -20.8151
10.9571 | -34.4099 | -20.8252
10.9633 | -34.4104 | -20.8356
10.9694 | -34.411 | -20.846
10.9755 | -34.4115 | -20.8559
10.9816 | -34.4121 | -20.866
10.9877 | -34.4127 | -20.8761
10.9938 | -34.4132 | -20.8864
10.9999 | -34.4138 | -20.8965
11.006 | -34.4144 | -20.9065
11.0121 | -34.4149 | -20.9169
11.0182 | -34.4155 | -20.9267
11.0243 | -34.416 | -20.9369
11.0304 | -34.4165 | -20.9469
11.0365 | -34.4171 | -20.9569
11.0426 | -34.4177 | -20.9667
11.0487 | -34.4182 | -20.9769
11.0548 | -34.4188 | -20.9869
11.0609 | -34.4194 | -20.9971
11.067 | -34.4199 | -21.0074
11.0731 | -34.4205 | -21.0176
11.0792 | -34.4211 | -21.0275
11.0853 | -34.4216 | -21.0373
11.0914 | -34.4222 | -21.0473
11.0975 | -34.4228 | -21.0576
11.1037 | -34.4233 | -21.0675
11.1098 | -34.4239 | -21.0778
11.1159 | -34.4245 | -21.0878
11.122 | -34.425 | -21.098
11.1281 | -34.4256 | -21.1078
11.1342 | -34.4261 | -21.1176
11.1403 | -34.4267 | -21.128
11.1464 | -34.4272 | -21.138
11.1525 | -34.4278 | -21.148
11.1586 | -34.4284 | -21.1586
11.1647 | -34.429 | -21.1685
11.1708 | -34.4295 | -21.1784
11.1769 | -34.43 | -21.188
11.183 | -34.4306 | -21.1978
11.1891 | -34.4312 | -21.2082
11.1952 | -34.4317 | -21.2186
11.2013 | -34.4323 | -21.2286
11.2074 | -34.4328 | -21.2384
11.2135 | -34.4333 | -21.2484
11.2196 | -34.4339 | -21.2583
11.2257 | -34.4344 | -21.2682
11.2318 | -34.4349 | -21.2784
11.2379 | -34.4354 | -21.2882
11.244 | -34.4359 | -21.2982
11.2502 | -34.4365 | -21.3082
11.2563 | -34.4371 | -21.3182
11.2624 | -34.4376 | -21.3282
11.2685 | -34.4382 | -21.338
11.2746 | -34.4387 | -21.3477
11.2807 | -34.4392 | -21.3582
11.2868 | -34.4398 | -21.3684
11.2929 | -34.4404 | -21.378
11.299 | -34.4409 | -21.3882
11.3051 | -34.4415 | -21.3982
11.3112 | -34.442 | -21.4082
11.3173 | -34.4426 | -21.4183
11.3234 | -34.4431 | -21.4283
11.3295 | -34.4437 | -21.4381
11.3356 | -34.4443 | -21.448
11.3417 | -34.4448 | -21.4581
11.3478 | -34.4453 | -21.468
11.3539 | -34.4459 | -21.4778
11.36 | -34.4464 | -21.4875
11.3661 | -34.447 | -21.4974
11.3722 | -34.4475 | -21.5073
11.3783 | -34.4481 | -21.5172
11.3844 | -34.4486 | -21.5272
11.3906 | -34.4491 | -21.5374
11.3967 | -34.4497 | -21.5471
11.4028 | -34.4502 | -21.5571
11.4089 | -34.4508 | -21.5668
11.415 | -34.4513 | -21.5767
11.4211 | -34.4519 | -21.5865
11.4272 | -34.4524 | -21.5964
11.4333 | -34.4529 | -21.6062
11.4394 | -34.4535 | -21.6164
11.4455 | -34.454 | -21.6263
11.4516 | -34.4546 | -21.6362
11.4577 | -34.4551 | -21.6464
11.4638 | -34.4557 | -21.6557
11.4699 | -34.4562 | -21.666
11.476 | -34.4568 | -21.6763
11.4821 | -34.4574 | -21.686
11.4882 | -34.4579 | -21.6962
11.4943 | -34.4585 | -21.706
11.5004 | -34.4591 | -21.7156
11.5065 | -34.4596 | -21.7254
11.5126 | -34.4602 | -21.735
11.5187 | -34.4608 | -21.7448
11.5248 | -34.4613 | -21.7549
11.5309 | -34.4619 | -21.7647
11.5371 | -34.4624 | -21.7745
11.5432 | -34.4629 | -21.7845
11.5493 | -34.4635 | -21.7945
11.5554 | -34.464 | -21.8043
11.5615 | -34.4646 | -21.8144
11.5676 | -34.4651 | -21.8245
11.5737 | -34.4657 | -21.8341
11.5798 | -34.4663 | -21.8437
11.5859 | -34.4669 | -21.8534
11.592 | -34.4674 | -21.8633
11.5981 | -34.468 | -21.873
11.6042 | -34.4686 | -21.883
11.6103 | -34.4691 | -21.8927
11.6164 | -34.4697 | -21.9024
11.6225 | -34.4702 | -21.9122
11.6286 | -34.4708 | -21.9222
11.6347 | -34.4713 | -21.9323
11.6408 | -34.4719 | -21.942
11.6469 | -34.4725 | -21.952
11.653 | -34.473 | -21.9616
11.6591 | -34.4736 | -21.9715
11.6652 | -34.4741 | -21.9815
11.6713 | -34.4747 | -21.9913
11.6775 | -34.4753 | -22.001
11.6836 | -34.4758 | -22.011
11.6897 | -34.4764 | -22.0206
11.6958 | -34.477 | -22.0308
11.7019 | -34.4775 | -22.0406
11.708 | -34.4781 | -22.0506
11.7141 | -34.4787 | -22.0604
11.7202 | -34.4793 | -22.07
11.7263 | -34.4798 | -22.08
11.7324 | -34.4804 | -22.09
11.7385 | -34.4809 | -22.1
11.7446 | -34.4814 | -22.1099
11.7507 | -34.482 | -22.1196
11.7568 | -34.4826 | -22.1295
11.7629 | -34.4831 | -22.1393
11.769 | -34.4837 | -22.1492
11.7751 | -34.4843 | -22.1587
11.7812 | -34.4848 | -22.1683
11.7873 | -34.4854 | -22.1785
11.7934 | -34.486 | -22.1883
11.7995 | -34.4865 | -22.1978
11.8056 | -34.4871 | -22.2079
11.8117 | -34.4876 | -22.2178
11.8178 | -34.4882 | -22.2275
11.824 | -34.4888 | -22.2372
11.8301 | -34.4894 | -22.2466
11.8362 | -34.4899 | -22.2561
11.8423 | -34.4904 | -22.266
11.8484 | -34.491 | -22.2757
11.8545 | -34.4916 | -22.2857
11.8606 | -34.4921 | -22.2955
11.8667 | -34.4927 | -22.3054
11.8728 | -34.4932 | -22.3152
11.8789 | -34.4938 | -22.3253
11.885 | -34.4943 | -22.3351
11.8911 | -34.4949 | -22.345
11.8972 | -34.4955 | -22.3545
11.9033 | -34.496 | -22.3638
11.9094 | -34.4965 | -22.3733
11.9155 | -34.4971 | -22.3831
11.9216 | -34.4977 | -22.3928
11.9277 | -34.4983 | -22.4023
11.9338 | -34.4988 | -22.4122
11.9399 | -34.4993 | -22.4222
11.946 | -34.4999 | -22.432
11.9521 | -34.5005 | -22.4416
11.9582 | -34.5011 | -22.4512
11.9644 | -34.5017 | -22.4608
11.9705 | -34.5022 | -22.4707
11.9766 | -34.5028 | -22.4804
11.9827 | -34.5033 | -22.4901
11.9888 | -34.5039 | -22.4996
11.9949 | -34.5044 | -22.5092
12.001 | -34.505 | -22.519
12.0071 | -34.5056 | -22.5285
12.0132 | -34.5061 | -22.5379
12.0193 | -34.5067 | -22.5479
12.0254 | -34.5073 | -22.5575
12.0315 | -34.5078 | -22.567
12.0376 | -34.5084 | -22.5765
12.0437 | -34.5089 | -22.586
12.0498 | -34.5095 | -22.596
12.0559 | -34.5101 | -22.6057
12.062 | -34.5106 | -22.6153
12.0681 | -34.5112 | -22.625
12.0742 | -34.5118 | -22.6349
12.0803 | -34.5123 | -22.6445
12.0864 | -34.5129 | -22.6542
12.0925 | -34.5135 | -22.6639
12.0986 | -34.5141 | -22.6733
12.1047 | -34.5147 | -22.6829
12.1109 | -34.5152 | -22.6927
12.117 | -34.5158 | -22.7022
12.1231 | -34.5164 | -22.7119
12.1292 | -34.517 | -22.7217
12.1353 | -34.5176 | -22.7311
12.1414 | -34.5181 | -22.7409
12.1475 | -34.5187 | -22.7507
12.1536 | -34.5192 | -22.7604
12.1597 | -34.5198 | -22.7701
12.1658 | -34.5203 | -22.7798
12.1719 | -34.521 | -22.7895
12.178 | -34.5215 | -22.7993
12.1841 | -34.5221 | -22.8091
12.1902 | -34.5227 | -22.8188
12.1963 | -34.5233 | -22.8285
12.2024 | -34.5239 | -22.838
12.2085 | -34.5244 | -22.8477
12.2146 | -34.525 | -22.8573
12.2207 | -34.5255 | -22.8669
12.2268 | -34.5261 | -22.8765
12.2329 | -34.5267 | -22.8862
12.239 | -34.5272 | -22.8956
12.2451 | -34.5278 | -22.905
12.2513 | -34.5284 | -22.9148
12.2574 | -34.529 | -22.9247
12.2635 | -34.5295 | -22.9339
12.2696 | -34.5301 | -22.9434
12.2757 | -34.5306 | -22.9527
12.2818 | -34.5312 | -22.9621
12.2879 | -34.5317 | -22.9716
12.294 | -34.5323 | -22.9811
12.3001 | -34.5329 | -22.9909
12.3062 | -34.5335 | -23.0005
12.3123 | -34.5341 | -23.0101
12.3184 | -34.5346 | -23.0195
12.3245 | -34.5352 | -23.0292
12.3306 | -34.5358 | -23.0385
12.3367 | -34.5364 | -23.048
12.3428 | -34.537 | -23.0574
12.3489 | -34.5375 | -23.0665
12.355 | -34.5381 | -23.0761
12.3611 | -34.5386 | -23.0857
12.3672 | -34.5392 | -23.0951
12.3733 | -34.5398 | -23.105
12.3794 | -34.5403 | -23.1141
12.3855 | -34.5409 | -23.1235
12.3916 | -34.5414 | -23.1332
12.3978 | -34.542 | -23.143
12.4039 | -34.5425 | -23.1525
12.41 | -34.5432 | -23.1622
12.4161 | -34.5438 | -23.1715
12.4222 | -34.5443 | -23.181
12.4283 | -34.5449 | -23.1905
12.4344 | -34.5455 | -23.2001
12.4405 | -34.546 | -23.2097
12.4466 | -34.5466 | -23.2191
12.4527 | -34.5472 | -23.2287
12.4588 | -34.5478 | -23.2382
12.4649 | -34.5483 | -23.248
12.471 | -34.5489 | -23.2579
12.4771 | -34.5495 | -23.2669
12.4832 | -34.55 | -23.2764
12.4893 | -34.5506 | -23.2857
12.4954 | -34.5513 | -23.2951
12.5015 | -34.5519 | -23.3044
12.5076 | -34.5524 | -23.314
12.5137 | -34.5529 | -23.3239
12.5198 | -34.5535 | -23.3331
12.5259 | -34.5541 | -23.3425
12.532 | -34.5547 | -23.3522
12.5382 | -34.5553 | -23.3616
12.5443 | -34.5558 | -23.3711
12.5504 | -34.5564 | -23.3807
12.5565 | -34.557 | -23.39
12.5626 | -34.5576 | -23.3997
12.5687 | -34.5582 | -23.4094
12.5748 | -34.5587 | -23.4188
12.5809 | -34.5593 | -23.428
12.587 | -34.5598 | -23.4378
12.5931 | -34.5605 | -23.4473
12.5992 | -34.561 | -23.457
12.6053 | -34.5616 | -23.4663
12.6114 | -34.5622 | -23.4758
12.6175 | -34.5628 | -23.4855
12.6236 | -34.5633 | -23.4947
12.6297 | -34.5639 | -23.504
12.6358 | -34.5645 | -23.5133
12.6419 | -34.5652 | -23.523
12.648 | -34.5657 | -23.5327
12.6541 | -34.5663 | -23.5421
12.6602 | -34.5669 | -23.5512
12.6663 | -34.5674 | -23.5605
12.6724 | -34.568 | -23.57
12.6785 | -34.5686 | -23.5797
12.6847 | -34.5692 | -23.5892
12.6908 | -34.5698 | -23.5986
12.6969 | -34.5704 | -23.608
12.703 | -34.571 | -23.6176
12.7091 | -34.5715 | -23.6271
12.7152 | -34.5721 | -23.6365
12.7213 | -34.5727 | -23.6461
12.7274 | -34.5733 | -23.6556
12.7335 | -34.5738 | -23.6646
12.7396 | -34.5744 | -23.6737
12.7457 | -34.575 | -23.6831
12.7518 | -34.5756 | -23.6926
12.7579 | -34.5761 | -23.7023
12.764 | -34.5767 | -23.7118
12.7701 | -34.5772 | -23.7209
12.7762 | -34.5778 | -23.7304
12.7823 | -34.5783 | -23.7399
12.7884 | -34.5789 | -23.7493
12.7945 | -34.5795 | -23.7584
12.8006 | -34.5801 | -23.7673
12.8067 | -34.5806 | -23.7769
12.8128 | -34.5812 | -23.786
12.8189 | -34.5817 | -23.7951
12.8251 | -34.5822 | -23.8044
12.8312 | -34.5828 | -23.8138
12.8373 | -34.5834 | -23.8232
12.8434 | -34.584 | -23.8325
12.8495 | -34.5845 | -23.8418
12.8556 | -34.5851 | -23.8511
12.8617 | -34.5856 | -23.8603
12.8678 | -34.5862 | -23.8697
12.8739 | -34.5868 | -23.8787
12.88 | -34.5873 | -23.8881
12.8861 | -34.5879 | -23.8975
12.8922 | -34.5884 | -23.9069
12.8983 | -34.589 | -23.9159
12.9044 | -34.5896 | -23.9252
12.9105 | -34.5903 | -23.9345
12.9166 | -34.5908 | -23.9441
12.9227 | -34.5914 | -23.9539
12.9288 | -34.592 | -23.9633
12.9349 | -34.5926 | -23.973
12.941 | -34.5932 | -23.9824
12.9471 | -34.5939 | -23.9918
12.9532 | -34.5945 | -24.0011
12.9593 | -34.595 | -24.0107
12.9654 | -34.5956 | -24.0202
12.9716 | -34.5962 | -24.0296
12.9777 | -34.5967 | -24.0387
12.9838 | -34.5973 | -24.0479
12.9899 | -34.5979 | -24.0573
12.996 | -34.5984 | -24.0668
13.0021 | -34.5989 | -24.0762
13.0082 | -34.5995 | -24.0855
13.0143 | -34.6001 | -24.0945
13.0204 | -34.6006 | -24.1036
13.0265 | -34.6012 | -24.1127
13.0326 | -34.6018 | -24.1222
13.0387 | -34.6024 | -24.1318
13.0448 | -34.6029 | -24.141
13.0509 | -34.6035 | -24.1504
13.057 | -34.6042 | -24.1594
13.0631 | -34.6048 | -24.1685
13.0692 | -34.6054 | -24.1777
13.0753 | -34.606 | -24.187
13.0814 | -34.6065 | -24.1961
13.0875 | -34.6071 | -24.2051
13.0936 | -34.6077 | -24.2142
13.0997 | -34.6083 | -24.2232
13.1058 | -34.6088 | -24.232
13.112 | -34.6094 | -24.2412
13.1181 | -34.61 | -24.2505
13.1242 | -34.6105 | -24.2598
13.1303 | -34.6111 | -24.2693
13.1364 | -34.6117 | -24.2786
13.1425 | -34.6122 | -24.2878
13.1486 | -34.6128 | -24.2969
13.1547 | -34.6134 | -24.3059
13.1608 | -34.614 | -24.3151
13.1669 | -34.6146 | -24.3242
13.173 | -34.6151 | -24.3334
13.1791 | -34.6157 | -24.3427
13.1852 | -34.6162 | -24.3522
13.1913 | -34.6168 | -24.3613
13.1974 | -34.6173 | -24.3706
13.2035 | -34.6179 | -24.3797
13.2096 | -34.6185 | -24.389
13.2157 | -34.6191 | -24.3981
13.2218 | -34.6197 | -24.4075
13.2279 | -34.6203 | -24.4166
13.234 | -34.6209 | -24.4258
13.2401 | -34.6215 | -24.435
13.2462 | -34.622 | -24.4441
13.2524 | -34.6226 | -24.4533
13.2585 | -34.6232 | -24.4624
13.2646 | -34.6237 | -24.4718
13.2707 | -34.6243 | -24.4811
13.2768 | -34.6249 | -24.4906
13.2829 | -34.6255 | -24.4999
13.289 | -34.626 | -24.5091
13.2951 | -34.6266 | -24.5182
13.3012 | -34.6272 | -24.5274
13.3073 | -34.6277 | -24.5369
13.3134 | -34.6283 | -24.5461
13.3195 | -34.6289 | -24.5551
13.3256 | -34.6295 | -24.5644
13.3317 | -34.63 | -24.5735
13.3378 | -34.6306 | -24.5828
13.3439 | -34.6312 | -24.592
13.35 | -34.6318 | -24.6012
13.3561 | -34.6323 | -24.6105
13.3622 | -34.6328 | -24.6197
13.3683 | -34.6335 | -24.6287
13.3744 | -34.634 | -24.6378
13.3805 | -34.6346 | -24.6471
13.3866 | -34.6351 | -24.656
13.3927 | -34.6357 | -24.6651
13.3989 | -34.6362 | -24.6741
13.405 | -34.6369 | -24.6832
13.4111 | -34.6375 | -24.6924
13.4172 | -34.638 | -24.7013
13.4233 | -34.6386 | -24.7107
13.4294 | -34.6391 | -24.7196
13.4355 | -34.6397 | -24.7292
13.4416 | -34.6403 | -24.7384
13.4477 | -34.6408 | -24.7476
13.4538 | -34.6414 | -24.7568
13.4599 | -34.642 | -24.7661
13.466 | -34.6426 | -24.7755
13.4721 | -34.6432 | -24.7848
13.4782 | -34.6437 | -24.7938
13.4843 | -34.6444 | -24.803
13.4904 | -34.6449 | -24.8117
13.4965 | -34.6455 | -24.8212
13.5026 | -34.6461 | -24.8303
13.5087 | -34.6467 | -24.8391
13.5148 | -34.6472 | -24.8485
13.5209 | -34.6478 | -24.8574
13.527 | -34.6483 | -24.8669
13.5331 | -34.6489 | -24.876
13.5393 | -34.6495 | -24.8851
13.5454 | -34.65 | -24.8943
13.5515 | -34.6506 | -24.9035
13.5576 | -34.6512 | -24.9127
13.5637 | -34.6518 | -24.9218
13.5698 | -34.6524 | -24.9311
13.5759 | -34.6529 | -24.9399
13.582 | -34.6534 | -24.9491
13.5881 | -34.654 | -24.9582
13.5942 | -34.6546 | -24.9674
13.6003 | -34.6552 | -24.9766
13.6064 | -34.6557 | -24.9859
13.6125 | -34.6563 | -24.9953
13.6186 | -34.6569 | -25.0044
13.6247 | -34.6575 | -25.0136
13.6308 | -34.658 | -25.0224
13.6369 | -34.6586 | -25.0316
13.643 | -34.6592 | -25.0405
13.6491 | -34.6598 | -25.0496
13.6552 | -34.6604 | -25.0586
13.6613 | -34.661 | -25.0679
13.6674 | -34.6615 | -25.0771
13.6735 | -34.6621 | -25.0864
13.6796 | -34.6627 | -25.0954
13.6858 | -34.6632 | -25.1045
13.6919 | -34.6638 | -25.1138
13.698 | -34.6644 | -25.1229
13.7041 | -34.6649 | -25.1323
13.7102 | -34.6655 | -25.1415
13.7163 | -34.6661 | -25.1507
13.7224 | -34.6667 | -25.1597
13.7285 | -34.6673 | -25.169
13.7346 | -34.6678 | -25.1781
13.7407 | -34.6684 | -25.1872
13.7468 | -34.669 | -25.1964
13.7529 | -34.6695 | -25.2053
13.759 | -34.6701 | -25.2146
13.7651 | -34.6707 | -25.2238
13.7712 | -34.6712 | -25.2331
13.7773 | -34.6718 | -25.2423
13.7834 | -34.6724 | -25.2514
13.7895 | -34.673 | -25.2604
13.7956 | -34.6736 | -25.2697
13.8017 | -34.6742 | -25.2786
13.8078 | -34.6748 | -25.2878
13.8139 | -34.6754 | -25.2968
13.82 | -34.676 | -25.3057
13.8262 | -34.6765 | -25.3151
13.8323 | -34.6771 | -25.3245
13.8384 | -34.6777 | -25.3338
13.8445 | -34.6783 | -25.343
13.8506 | -34.6789 | -25.3523
13.8567 | -34.6794 | -25.3616
13.8628 | -34.68 | -25.3708
13.8689 | -34.6805 | -25.3801
13.875 | -34.6812 | -25.3892
13.8811 | -34.6818 | -25.3982
13.8872 | -34.6824 | -25.4073
13.8933 | -34.683 | -25.4162
13.8994 | -34.6835 | -25.4251
13.9055 | -34.6841 | -25.4342
13.9116 | -34.6847 | -25.443
13.9177 | -34.6853 | -25.4519
13.9238 | -34.6859 | -25.4611
13.9299 | -34.6865 | -25.4702
13.936 | -34.6871 | -25.4792
13.9421 | -34.6877 | -25.4883
13.9482 | -34.6883 | -25.4974
13.9543 | -34.6888 | -25.5063
13.9604 | -34.6894 | -25.5156
13.9665 | -34.6901 | -25.5246
13.9727 | -34.6906 | -25.5336
13.9788 | -34.6912 | -25.5429
13.9849 | -34.6917 | -25.5521
13.991 | -34.6923 | -25.5615
13.9971 | -34.6928 | -25.5708
14.0032 | -34.6934 | -25.5801
14.0093 | -34.6939 | -25.589
14.0154 | -34.6945 | -25.5983
14.0215 | -34.6951 | -25.6072
14.0276 | -34.6957 | -25.6162
14.0337 | -34.6963 | -25.6254
14.0398 | -34.6969 | -25.6343
14.0459 | -34.6974 | -25.6435
14.052 | -34.698 | -25.6527
14.0581 | -34.6985 | -25.6615
14.0642 | -34.6991 | -25.6707
14.0703 | -34.6997 | -25.6798
14.0764 | -34.7003 | -25.6893
14.0825 | -34.7009 | -25.6984
14.0886 | -34.7014 | -25.7071
14.0947 | -34.702 | -25.716
14.1008 | -34.7026 | -25.7253
14.1069 | -34.7032 | -25.7343
14.1131 | -34.7037 | -25.7436
14.1192 | -34.7043 | -25.7524
14.1253 | -34.7048 | -25.7612
14.1314 | -34.7054 | -25.7703
14.1375 | -34.706 | -25.7795
14.1436 | -34.7067 | -25.7885
14.1497 | -34.7073 | -25.7974
14.1558 | -34.7078 | -25.8066
14.1619 | -34.7084 | -25.8156
14.168 | -34.709 | -25.8246
14.1741 | -34.7096 | -25.8334
14.1802 | -34.7102 | -25.8424
14.1863 | -34.7108 | -25.8515
14.1924 | -34.7114 | -25.8606
14.1985 | -34.712 | -25.8695
14.2046 | -34.7125 | -25.8784
14.2107 | -34.7131 | -25.8871
14.2168 | -34.7137 | -25.8965
14.2229 | -34.7143 | -25.9056
14.229 | -34.7149 | -25.9147
14.2351 | -34.7155 | -25.9235
14.2412 | -34.716 | -25.9325
14.2473 | -34.7167 | -25.9415
14.2534 | -34.7173 | -25.9504
14.2596 | -34.7178 | -25.9595
14.2657 | -34.7184 | -25.9686
14.2718 | -34.719 | -25.9778
14.2779 | -34.7196 | -25.9866
14.284 | -34.7202 | -25.9956
14.2901 | -34.7208 | -26.0044
14.2962 | -34.7214 | -26.0132
14.3023 | -34.722 | -26.0223
14.3084 | -34.7226 | -26.0315
14.3145 | -34.7232 | -26.0405
14.3206 | -34.7237 | -26.0496
14.3267 | -34.7243 | -26.0586
14.3328 | -34.7249 | -26.0675
14.3389 | -34.7255 | -26.0763
14.345 | -34.726 | -26.0855
14.3511 | -34.7266 | -26.0942
14.3572 | -34.7272 | -26.1035
14.3633 | -34.7278 | -26.1124
14.3694 | -34.7283 | -26.1213
14.3755 | -34.7289 | -26.1301
14.3816 | -34.7295 | -26.139
14.3877 | -34.7301 | -26.1478
14.3938 | -34.7307 | -26.157
14.4 | -34.7313 | -26.1658
14.4061 | -34.7318 | -26.1748
14.4122 | -34.7324 | -26.1839
14.4183 | -34.7329 | -26.1928
14.4244 | -34.7335 | -26.2017
14.4305 | -34.7341 | -26.2109
14.4366 | -34.7347 | -26.22
14.4427 | -34.7353 | -26.2291
14.4488 | -34.7358 | -26.2379
14.4549 | -34.7365 | -26.2467
14.461 | -34.7371 | -26.2557
14.4671 | -34.7376 | -26.2648
14.4732 | -34.7382 | -26.2734
14.4793 | -34.7388 | -26.283
14.4854 | -34.7395 | -26.2916
14.4915 | -34.74 | -26.3003
14.4976 | -34.7407 | -26.3094
14.5037 | -34.7413 | -26.3185
14.5098 | -34.7419 | -26.3273
14.5159 | -34.7425 | -26.3364
14.522 | -34.743 | -26.3455
14.5281 | -34.7437 | -26.3542
14.5342 | -34.7442 | -26.3633
14.5403 | -34.7448 | -26.3723
14.5465 | -34.7454 | -26.3816
14.5526 | -34.746 | -26.3908
14.5587 | -34.7466 | -26.3998
14.5648 | -34.7472 | -26.4085
14.5709 | -34.7478 | -26.4176
14.577 | -34.7484 | -26.4266
14.5831 | -34.7489 | -26.4356
14.5892 | -34.7496 | -26.4442
14.5953 | -34.7502 | -26.4528
14.6014 | -34.7508 | -26.4617
14.6075 | -34.7514 | -26.4705
14.6136 | -34.7519 | -26.4792
14.6197 | -34.7525 | -26.4883
14.6258 | -34.7531 | -26.4973
14.6319 | -34.7537 | -26.5059
14.638 | -34.7543 | -26.5149
14.6441 | -34.7549 | -26.5234
14.6502 | -34.7555 | -26.5321
14.6563 | -34.7561 | -26.541
14.6624 | -34.7567 | -26.5499
14.6685 | -34.7573 | -26.5584
14.6746 | -34.7579 | -26.5674
14.6807 | -34.7584 | -26.5763
14.6869 | -34.759 | -26.5851
14.693 | -34.7596 | -26.5939
14.6991 | -34.7602 | -26.6026
14.7052 | -34.7608 | -26.6113
14.7113 | -34.7614 | -26.6204
14.7174 | -34.762 | -26.6292
14.7235 | -34.7626 | -26.6379
14.7296 | -34.7632 | -26.6468
14.7357 | -34.7638 | -26.6554
14.7418 | -34.7644 | -26.664
14.7479 | -34.765 | -26.673
14.754 | -34.7656 | -26.6817
14.7601 | -34.7662 | -26.6906
14.7662 | -34.7668 | -26.6995
14.7723 | -34.7674 | -26.7082
14.7784 | -34.768 | -26.7172
14.7845 | -34.7686 | -26.7261
14.7906 | -34.7692 | -26.735
14.7967 | -34.7698 | -26.7438
14.8028 | -34.7703 | -26.7525
14.8089 | -34.7709 | -26.7611
14.815 | -34.7715 | -26.7697
14.8211 | -34.7721 | -26.7785
14.8272 | -34.7727 | -26.7875
14.8334 | -34.7733 | -26.7963
14.8395 | -34.7739 | -26.8051
14.8456 | -34.7745 | -26.8137
14.8517 | -34.7751 | -26.8227
14.8578 | -34.7757 | -26.8315
14.8639 | -34.7763 | -26.8404
14.87 | -34.7769 | -26.8493
14.8761 | -34.7775 | -26.8578
14.8822 | -34.7781 | -26.8667
14.8883 | -34.7787 | -26.8753
14.8944 | -34.7793 | -26.8841
14.9005 | -34.7799 | -26.8927
14.9066 | -34.7805 | -26.9016
14.9127 | -34.7811 | -26.9102
14.9188 | -34.7817 | -26.9192
14.9249 | -34.7823 | -26.9279
14.931 | -34.7829 | -26.9366
14.9371 | -34.7835 | -26.9453
14.9432 | -34.7841 | -26.9543
14.9493 | -34.7847 | -26.963
14.9554 | -34.7853 | -26.9717
14.9615 | -34.7859 | -26.9806
14.9676 | -34.7865 | -26.9896
14.9738 | -34.7871 | -26.9982
14.9799 | -34.7877 | -27.0073
14.986 | -34.7883 | -27.0157
14.9921 | -34.7889 | -27.0243
14.9982 | -34.7895 | -27.0333
15.0043 | -34.7901 | -27.042
15.0104 | -34.7906 | -27.0507
15.0165 | -34.7913 | -27.0596
15.0226 | -34.7919 | -27.068
15.0287 | -34.7924 | -27.0771
15.0348 | -34.7931 | -27.0861
15.0409 | -34.7937 | -27.0951
15.047 | -34.7943 | -27.1041
15.0531 | -34.795 | -27.1129
15.0592 | -34.7956 | -27.1217
15.0653 | -34.7962 | -27.1305
15.0714 | -34.7968 | -27.139
15.0775 | -34.7974 | -27.1475
15.0836 | -34.798 | -27.1563
15.0897 | -34.7987 | -27.1649
15.0958 | -34.7993 | -27.1734
15.1019 | -34.7998 | -27.1822
15.108 | -34.8004 | -27.191
15.1141 | -34.801 | -27.1997
15.1203 | -34.8017 | -27.2086
15.1264 | -34.8024 | -27.2174
15.1325 | -34.803 | -27.226
15.1386 | -34.8036 | -27.2349
15.1447 | -34.8042 | -27.2437
15.1508 | -34.8049 | -27.2527
15.1569 | -34.8055 | -27.2617
15.163 | -34.8062 | -27.27
15.1691 | -34.8068 | -27.2786
15.1752 | -34.8074 | -27.2872
15.1813 | -34.808 | -27.296
15.1874 | -34.8086 | -27.3048
15.1935 | -34.8093 | -27.3132
15.1996 | -34.8099 | -27.3221
15.2057 | -34.8105 | -27.3307
15.2118 | -34.8111 | -27.3394
15.2179 | -34.8118 | -27.3479
15.224 | -34.8123 | -27.3567
15.2301 | -34.813 | -27.3651
15.2362 | -34.8136 | -27.3738
15.2423 | -34.8142 | -27.3827
15.2484 | -34.8148 | -27.3913
15.2545 | -34.8155 | -27.3999
15.2607 | -34.8161 | -27.4087
15.2668 | -34.8168 | -27.4171
15.2729 | -34.8174 | -27.4257
15.279 | -34.818 | -27.4345
15.2851 | -34.8186 | -27.4431
15.2912 | -34.8192 | -27.4517
15.2973 | -34.8198 | -27.4603
15.3034 | -34.8205 | -27.4687
15.3095 | -34.8211 | -27.4772
15.3156 | -34.8217 | -27.4861
15.3217 | -34.8223 | -27.4945
15.3278 | -34.823 | -27.5032
15.3339 | -34.8236 | -27.5116
15.34 | -34.8242 | -27.5201
15.3461 | -34.8248 | -27.5288
15.3522 | -34.8254 | -27.5372
15.3583 | -34.8261 | -27.5459
15.3644 | -34.8267 | -27.5546
15.3705 | -34.8273 | -27.5629
15.3766 | -34.828 | -27.5718
15.3827 | -34.8286 | -27.5802
15.3888 | -34.8292 | -27.5885
15.3949 | -34.8298 | -27.597
15.401 | -34.8304 | -27.6057
15.4072 | -34.8311 | -27.6139
15.4133 | -34.8317 | -27.6226
15.4194 | -34.8323 | -27.6313
15.4255 | -34.833 | -27.6399
15.4316 | -34.8336 | -27.6486
15.4377 | -34.8342 | -27.6571
15.4438 | -34.8349 | -27.6653
15.4499 | -34.8355 | -27.6739
15.456 | -34.8361 | -27.6824
15.4621 | -34.8367 | -27.691
15.4682 | -34.8374 | -27.6997
15.4743 | -34.838 | -27.7082
15.4804 | -34.8387 | -27.7168
15.4865 | -34.8393 | -27.7253
15.4926 | -34.8399 | -27.7339
15.4987 | -34.8405 | -27.7425
15.5048 | -34.8411 | -27.751
15.5109 | -34.8418 | -27.7592
15.517 | -34.8424 | -27.7679
15.5231 | -34.8431 | -27.7765
15.5292 | -34.8437 | -27.785
15.5353 | -34.8443 | -27.7935
15.5414 | -34.845 | -27.8022
15.5476 | -34.8456 | -27.8108
15.5537 | -34.8462 | -27.8193
15.5598 | -34.8469 | -27.8278
15.5659 | -34.8475 | -27.8366
15.572 | -34.8481 | -27.8448
15.5781 | -34.8488 | -27.8534
15.5842 | -34.8494 | -27.8616
15.5903 | -34.8501 | -27.8699
15.5964 | -34.8507 | -27.8783
15.6025 | -34.8513 | -27.8867
15.6086 | -34.852 | -27.8952
15.6147 | -34.8526 | -27.9038
15.6208 | -34.8532 | -27.912
15.6269 | -34.8539 | -27.9204
15.633 | -34.8545 | -27.9289
15.6391 | -34.8551 | -27.9375
15.6452 | -34.8558 | -27.9459
15.6513 | -34.8564 | -27.9545
15.6574 | -34.8571 | -27.9627
15.6635 | -34.8577 | -27.9711
15.6696 | -34.8583 | -27.9795
15.6757 | -34.859 | -27.9883
15.6818 | -34.8596 | -27.997
15.688 | -34.8603 | -28.0054
15.6941 | -34.8609 | -28.0139
15.7002 | -34.8616 | -28.0223
15.7063 | -34.8622 | -28.0309
15.7124 | -34.8628 | -28.0395
15.7185 | -34.8634 | -28.0481
15.7246 | -34.8641 | -28.0563
15.7307 | -34.8647 | -28.0649
15.7368 | -34.8654 | -28.0735
15.7429 | -34.866 | -28.0819
15.749 | -34.8666 | -28.0904
15.7551 | -34.8672 | -28.0989
15.7612 | -34.8679 | -28.1068
15.7673 | -34.8685 | -28.1152
15.7734 | -34.8691 | -28.1234
15.7795 | -34.8697 | -28.1322
15.7856 | -34.8704 | -28.1405
15.7917 | -34.871 | -28.1487
15.7978 | -34.8717 | -28.1572
15.8039 | -34.8723 | -28.1655
15.81 | -34.8729 | -28.1739
15.8161 | -34.8735 | -28.1825
15.8222 | -34.8742 | -28.1909
15.8283 | -34.8749 | -28.1991
15.8345 | -34.8755 | -28.2076
15.8406 | -34.8762 | -28.216
15.8467 | -34.8768 | -28.2245
15.8528 | -34.8774 | -28.2328
15.8589 | -34.8781 | -28.2411
15.865 | -34.8787 | -28.2493
15.8711 | -34.8794 | -28.258
15.8772 | -34.8801 | -28.2665
15.8833 | -34.8807 | -28.2752
15.8894 | -34.8813 | -28.2836
15.8955 | -34.882 | -28.2921
15.9016 | -34.8827 | -28.3004
15.9077 | -34.8833 | -28.3089
15.9138 | -34.8839 | -28.3171
15.9199 | -34.8846 | -28.3256
15.926 | -34.8852 | -28.3342
15.9321 | -34.8859 | -28.3425
15.9382 | -34.8866 | -28.3509
15.9443 | -34.8872 | -28.3596
15.9504 | -34.8879 | -28.3679
15.9565 | -34.8886 | -28.3763
15.9626 | -34.8892 | -28.3848
15.9687 | -34.8899 | -28.393
15.9749 | -34.8905 | -28.4015
15.981 | -34.8912 | -28.41
15.9871 | -34.8918 | -28.4184
15.9932 | -34.8925 | -28.4268
15.9993 | -34.8931 | -28.435
16.0054 | -34.8938 | -28.4431
16.0115 | -34.8945 | -28.4515
16.0176 | -34.8951 | -28.4601
16.0237 | -34.8957 | -28.4688
16.0298 | -34.8964 | -28.4767
16.0359 | -34.897 | -28.485
16.042 | -34.8977 | -28.4934
16.0481 | -34.8984 | -28.5018
16.0542 | -34.899 | -28.5104
16.0603 | -34.8997 | -28.5187
16.0664 | -34.9003 | -28.5272
16.0725 | -34.901 | -28.5353
16.0786 | -34.9017 | -28.5439
16.0847 | -34.9023 | -28.552
16.0908 | -34.9029 | -28.5604
16.0969 | -34.9036 | -28.5686
16.103 | -34.9042 | -28.5773
16.1091 | -34.9049 | -28.5855
16.1152 | -34.9055 | -28.5939
16.1214 | -34.9062 | -28.6024
16.1275 | -34.9068 | -28.6107
16.1336 | -34.9075 | -28.6191
16.1397 | -34.9082 | -28.6275
16.1458 | -34.9088 | -28.6357
16.1519 | -34.9095 | -28.6443
16.158 | -34.9101 | -28.6528
16.1641 | -34.9108 | -28.6612
16.1702 | -34.9114 | -28.6697
16.1763 | -34.9121 | -28.6781
16.1824 | -34.9128 | -28.6862
16.1885 | -34.9134 | -28.6945
16.1946 | -34.9141 | -28.7028
16.2007 | -34.9147 | -28.7115
16.2068 | -34.9154 | -28.7198
16.2129 | -34.916 | -28.7283
16.219 | -34.9166 | -28.7365
16.2251 | -34.9173 | -28.7451
16.2312 | -34.9179 | -28.7534
16.2373 | -34.9185 | -28.7618
16.2434 | -34.9192 | -28.7702
16.2495 | -34.9199 | -28.7783
16.2556 | -34.9206 | -28.7867
16.2618 | -34.9212 | -28.7947
16.2679 | -34.9219 | -28.803
16.274 | -34.9225 | -28.8114
16.2801 | -34.9232 | -28.8195
16.2862 | -34.9239 | -28.8278
16.2923 | -34.9245 | -28.8363
16.2984 | -34.9251 | -28.8447
16.3045 | -34.9258 | -28.853
16.3106 | -34.9265 | -28.861
16.3167 | -34.9271 | -28.8694
16.3228 | -34.9278 | -28.8778
16.3289 | -34.9284 | -28.8861
16.335 | -34.9291 | -28.8946
16.3411 | -34.9297 | -28.9028
16.3472 | -34.9304 | -28.9113
16.3533 | -34.931 | -28.9198
16.3594 | -34.9317 | -28.9278
16.3655 | -34.9323 | -28.9363
16.3716 | -34.933 | -28.9445
16.3777 | -34.9336 | -28.9527
16.3838 | -34.9342 | -28.9609
16.3899 | -34.9349 | -28.9692
16.396 | -34.9356 | -28.9776
16.4021 | -34.9362 | -28.9858
16.4083 | -34.9369 | -28.9941
16.4144 | -34.9376 | -29.0023
16.4205 | -34.9382 | -29.0107
16.4266 | -34.9389 | -29.0189
16.4327 | -34.9396 | -29.027
16.4388 | -34.9402 | -29.0354
16.4449 | -34.9409 | -29.0437
16.451 | -34.9415 | -29.0521
16.4571 | -34.9421 | -29.0606
16.4632 | -34.9428 | -29.0686
16.4693 | -34.9434 | -29.077
16.4754 | -34.944 | -29.0854
16.4815 | -34.9447 | -29.0938
16.4876 | -34.9453 | -29.102
16.4937 | -34.946 | -29.1101
16.4998 | -34.9467 | -29.1183
16.5059 | -34.9474 | -29.1266
16.512 | -34.948 | -29.1352
16.5181 | -34.9487 | -29.1436
16.5242 | -34.9493 | -29.152
16.5303 | -34.9499 | -29.1599
16.5364 | -34.9506 | -29.1681
16.5425 | -34.9512 | -29.1765
16.5487 | -34.9519 | -29.1846
16.5548 | -34.9525 | -29.1929
16.5609 | -34.9532 | -29.2014
16.567 | -34.9538 | -29.2096
16.5731 | -34.9545 | -29.2178
16.5792 | -34.9551 | -29.226
16.5853 | -34.9558 | -29.2342
16.5914 | -34.9564 | -29.2426
16.5975 | -34.9571 | -29.2509
16.6036 | -34.9577 | -29.2591
16.6097 | -34.9584 | -29.2673
16.6158 | -34.959 | -29.2757
16.6219 | -34.9597 | -29.2838
16.628 | -34.9603 | -29.2922
16.6341 | -34.961 | -29.3004
16.6402 | -34.9616 | -29.3083
16.6463 | -34.9623 | -29.3165
16.6524 | -34.9629 | -29.3246
16.6585 | -34.9635 | -29.333
16.6646 | -34.9642 | -29.3413
16.6707 | -34.9649 | -29.3493
16.6768 | -34.9655 | -29.3575
16.6829 | -34.9662 | -29.3658
16.689 | -34.9668 | -29.3741
16.6952 | -34.9675 | -29.3824
16.7013 | -34.9681 | -29.3908
16.7074 | -34.9688 | -29.3989
16.7135 | -34.9694 | -29.4071
16.7196 | -34.9701 | -29.4154
16.7257 | -34.9707 | -29.4235
16.7318 | -34.9713 | -29.4318
16.7379 | -34.972 | -29.44
16.744 | -34.9727 | -29.4479
16.7501 | -34.9733 | -29.4563
16.7562 | -34.974 | -29.4644
16.7623 | -34.9746 | -29.4727
16.7684 | -34.9753 | -29.4811
16.7745 | -34.976 | -29.489
16.7806 | -34.9766 | -29.4973
16.7867 | -34.9772 | -29.5054
16.7928 | -34.9779 | -29.5137
16.7989 | -34.9785 | -29.5218
16.805 | -34.9791 | -29.5301
16.8111 | -34.9798 | -29.5383
16.8172 | -34.9804 | -29.5469
16.8233 | -34.9811 | -29.5553
16.8294 | -34.9817 | -29.5635
16.8356 | -34.9824 | -29.572
16.8417 | -34.983 | -29.5802
16.8478 | -34.9836 | -29.5882
16.8539 | -34.9843 | -29.5963
16.86 | -34.9849 | -29.6046
16.8661 | -34.9856 | -29.6128
16.8722 | -34.9863 | -29.6211
16.8783 | -34.9869 | -29.6292
16.8844 | -34.9876 | -29.6372
16.8905 | -34.9882 | -29.6456
16.8966 | -34.9888 | -29.6539
16.9027 | -34.9895 | -29.6619
16.9088 | -34.9902 | -29.6706
16.9149 | -34.9908 | -29.6792
16.921 | -34.9915 | -29.6873
16.9271 | -34.9921 | -29.6957
16.9332 | -34.9928 | -29.704
16.9393 | -34.9934 | -29.7124
16.9454 | -34.994 | -29.7204
16.9515 | -34.9947 | -29.7288
16.9576 | -34.9953 | -29.7367
16.9637 | -34.996 | -29.745
16.9698 | -34.9967 | -29.7529
16.9759 | -34.9973 | -29.761
16.9821 | -34.998 | -29.7694
16.9882 | -34.9986 | -29.7776
16.9943 | -34.9993 | -29.7859
17.0004 | -34.9999 | -29.7941
17.0065 | -35.0006 | -29.8024
17.0126 | -35.0012 | -29.8105
17.0187 | -35.0019 | -29.8188
17.0248 | -35.0025 | -29.8266
17.0309 | -35.0031 | -29.8349
17.037 | -35.0038 | -29.8432
17.0431 | -35.0044 | -29.8512
17.0492 | -35.0051 | -29.8595
17.0553 | -35.0057 | -29.8678
17.0614 | -35.0063 | -29.8759
17.0675 | -35.0069 | -29.8843
17.0736 | -35.0076 | -29.8926
17.0797 | -35.0082 | -29.9004
17.0858 | -35.0088 | -29.9088
17.0919 | -35.0095 | -29.9172
17.098 | -35.0101 | -29.9252
17.1041 | -35.0108 | -29.9331
17.1102 | -35.0114 | -29.9415
17.1163 | -35.0121 | -29.9497
17.1225 | -35.0128 | -29.9577
17.1286 | -35.0134 | -29.9656
17.1347 | -35.014 | -29.9736
17.1408 | -35.0147 | -29.9819
17.1469 | -35.0153 | -29.9899
17.153 | -35.016 | -29.998
17.1591 | -35.0167 | -30.006
17.1652 | -35.0173 | -30.0141
17.1713 | -35.0179 | -30.0222
17.1774 | -35.0185 | -30.0305
17.1835 | -35.0192 | -30.0388
17.1896 | -35.0198 | -30.0468
17.1957 | -35.0204 | -30.055
17.2018 | -35.0211 | -30.0631
17.2079 | -35.0217 | -30.0712
17.214 | -35.0223 | -30.0793
17.2201 | -35.0229 | -30.0874
17.2262 | -35.0236 | -30.0957
17.2323 | -35.0242 | -30.1039
17.2384 | -35.0249 | -30.1118
17.2445 | -35.0255 | -30.1199
17.2506 | -35.0261 | -30.1281
17.2567 | -35.0267 | -30.1363
17.2628 | -35.0274 | -30.1447
17.269 | -35.028 | -30.1528
17.2751 | -35.0286 | -30.1609
17.2812 | -35.0292 | -30.169
17.2873 | -35.0299 | -30.1772
17.2934 | -35.0305 | -30.1853
17.2995 | -35.0311 | -30.1933
17.3056 | -35.0317 | -30.2013
17.3117 | -35.0324 | -30.2095
17.3178 | -35.033 | -30.2175
17.3239 | -35.0336 | -30.2255
17.33 | -35.0342 | -30.2335
17.3361 | -35.0348 | -30.2415
17.3422 | -35.0355 | -30.2495
17.3483 | -35.0361 | -30.2576
17.3544 | -35.0368 | -30.2656
17.3605 | -35.0374 | -30.2738
17.3666 | -35.0381 | -30.2821
17.3727 | -35.0387 | -30.2902
17.3788 | -35.0393 | -30.298
17.3849 | -35.0399 | -30.306
17.391 | -35.0406 | -30.314
17.3971 | -35.0413 | -30.3222
17.4032 | -35.0419 | -30.3306
17.4094 | -35.0425 | -30.3386
17.4155 | -35.0431 | -30.3466
17.4216 | -35.0437 | -30.3547
17.4277 | -35.0444 | -30.363
17.4338 | -35.045 | -30.3711
17.4399 | -35.0457 | -30.3793
17.446 | -35.0463 | -30.3872
17.4521 | -35.0469 | -30.3952
17.4582 | -35.0476 | -30.4031
17.4643 | -35.0482 | -30.4113
17.4704 | -35.0488 | -30.4192
17.4765 | -35.0494 | -30.4273
17.4826 | -35.0501 | -30.4353
17.4887 | -35.0507 | -30.4436
17.4948 | -35.0514 | -30.4516
17.5009 | -35.052 | -30.4595
17.507 | -35.0526 | -30.4678
17.5131 | -35.0532 | -30.4759
17.5192 | -35.0539 | -30.4837
17.5253 | -35.0545 | -30.4919
17.5314 | -35.0551 | -30.5
17.5375 | -35.0557 | -30.5082
17.5436 | -35.0563 | -30.5164
17.5497 | -35.0569 | -30.5247
17.5559 | -35.0575 | -30.5324
17.562 | -35.0582 | -30.5403
17.5681 | -35.0588 | -30.5483
17.5742 | -35.0594 | -30.5565
17.5803 | -35.0601 | -30.5644
17.5864 | -35.0607 | -30.5721
17.5925 | -35.0614 | -30.5804
17.5986 | -35.0621 | -30.5883
17.6047 | -35.0627 | -30.5967
17.6108 | -35.0633 | -30.6045
17.6169 | -35.064 | -30.6124
17.623 | -35.0646 | -30.6203
17.6291 | -35.0652 | -30.6283
17.6352 | -35.0658 | -30.6362
17.6413 | -35.0664 | -30.6443
17.6474 | -35.0671 | -30.6524
17.6535 | -35.0677 | -30.6605
17.6596 | -35.0683 | -30.6685
17.6657 | -35.069 | -30.6764
17.6718 | -35.0696 | -30.6846
17.6779 | -35.0702 | -30.6925
17.684 | -35.0708 | -30.7006
17.6901 | -35.0715 | -30.7085
17.6963 | -35.0721 | -30.7163
17.7024 | -35.0727 | -30.7243
17.7085 | -35.0733 | -30.7322
17.7146 | -35.074 | -30.7401
17.7207 | -35.0746 | -30.748
17.7268 | -35.0752 | -30.7559
17.7329 | -35.0758 | -30.7641
17.739 | -35.0765 | -30.7723
17.7451 | -35.0771 | -30.7803
17.7512 | -35.0778 | -30.7884
17.7573 | -35.0784 | -30.7964
17.7634 | -35.079 | -30.8041
17.7695 | -35.0796 | -30.8122
17.7756 | -35.0803 | -30.8202
17.7817 | -35.0809 | -30.828
17.7878 | -35.0815 | -30.836
17.7939 | -35.0821 | -30.844
17.8 | -35.0828 | -30.8517
17.8061 | -35.0834 | -30.8599
17.8122 | -35.084 | -30.8679
17.8183 | -35.0846 | -30.8759
17.8244 | -35.0852 | -30.8838
17.8305 | -35.0859 | -30.8919
17.8366 | -35.0865 | -30.8995
17.8428 | -35.0871 | -30.9077
17.8489 | -35.0878 | -30.9154
17.855 | -35.0884 | -30.9231
17.8611 | -35.089 | -30.931
17.8672 | -35.0896 | -30.9388
17.8733 | -35.0903 | -30.947
17.8794 | -35.0909 | -30.9551
17.8855 | -35.0915 | -30.963
17.8916 | -35.0922 | -30.9708
17.8977 | -35.0928 | -30.9788
17.9038 | -35.0934 | -30.9865
17.9099 | -35.094 | -30.9944
17.916 | -35.0946 | -31.0024
17.9221 | -35.0952 | -31.0104
17.9282 | -35.0958 | -31.0183
17.9343 | -35.0965 | -31.0264
17.9404 | -35.0971 | -31.0341
17.9465 | -35.0977 | -31.0422
17.9526 | -35.0983 | -31.0504
17.9587 | -35.099 | -31.0583
17.9648 | -35.0996 | -31.0662
17.9709 | -35.1002 | -31.0741
17.977 | -35.1008 | -31.0818
17.9832 | -35.1014 | -31.0898
17.9893 | -35.102 | -31.0979
17.9954 | -35.1027 | -31.106
18.0015 | -35.1033 | -31.114
18.0076 | -35.1039 | -31.1214
18.0137 | -35.1045 | -31.1296
18.0198 | -35.1051 | -31.1375
18.0259 | -35.1057 | -31.1456
18.032 | -35.1063 | -31.1533
18.0381 | -35.1069 | -31.1612
18.0442 | -35.1076 | -31.1691
18.0503 | -35.1082 | -31.1771
18.0564 | -35.1088 | -31.185
18.0625 | -35.1094 | -31.193
18.0686 | -35.11 | -31.2008
18.0747 | -35.1106 | -31.2087
18.0808 | -35.1112 | -31.2163
18.0869 | -35.1119 | -31.2243
18.093 | -35.1125 | -31.2322
18.0991 | -35.1131 | -31.2403
18.1052 | -35.1138 | -31.2484
18.1113 | -35.1144 | -31.256
18.1174 | -35.115 | -31.2636
18.1236 | -35.1157 | -31.2717
18.1297 | -35.1163 | -31.2795
18.1358 | -35.1169 | -31.2874
18.1419 | -35.1176 | -31.2954
18.148 | -35.1182 | -31.3034
18.1541 | -35.1189 | -31.311
18.1602 | -35.1195 | -31.3191
18.1663 | -35.1202 | -31.3269
18.1724 | -35.1208 | -31.3349
18.1785 | -35.1214 | -31.3429
18.1846 | -35.122 | -31.3506
18.1907 | -35.1226 | -31.3588
18.1968 | -35.1231 | -31.3668
18.2029 | -35.1238 | -31.3744
18.209 | -35.1244 | -31.3826
18.2151 | -35.1251 | -31.3907
18.2212 | -35.1257 | -31.3984
18.2273 | -35.1264 | -31.4062
18.2334 | -35.127 | -31.4141
18.2395 | -35.1275 | -31.4219
18.2456 | -35.1282 | -31.4298
18.2517 | -35.1288 | -31.4376
18.2578 | -35.1294 | -31.4452
18.2639 | -35.13 | -31.4533
18.2701 | -35.1307 | -31.4612
18.2762 | -35.1313 | -31.4692
18.2823 | -35.1319 | -31.4772
18.2884 | -35.1326 | -31.4851
18.2945 | -35.1332 | -31.4927
18.3006 | -35.1338 | -31.5006
18.3067 | -35.1344 | -31.5085
18.3128 | -35.1351 | -31.5163
18.3189 | -35.1357 | -31.5243
18.325 | -35.1363 | -31.5318
18.3311 | -35.1369 | -31.5397
18.3372 | -35.1375 | -31.5474
18.3433 | -35.1381 | -31.555
18.3494 | -35.1387 | -31.5629
18.3555 | -35.1394 | -31.5704
18.3616 | -35.14 | -31.5779
18.3677 | -35.1406 | -31.5855
18.3738 | -35.1412 | -31.5933
18.3799 | -35.1418 | -31.6009
18.386 | -35.1424 | -31.6087
18.3921 | -35.1431 | -31.6164
18.3982 | -35.1437 | -31.6239
18.4043 | -35.1442 | -31.6317
18.4105 | -35.1448 | -31.6396
18.4166 | -35.1455 | -31.6473
18.4227 | -35.1461 | -31.6552
18.4288 | -35.1467 | -31.6632
18.4349 | -35.1473 | -31.6709
18.441 | -35.1479 | -31.6787
18.4471 | -35.1485 | -31.6864
18.4532 | -35.1492 | -31.6943
18.4593 | -35.1498 | -31.702
18.4654 | -35.1504 | -31.7095
18.4715 | -35.151 | -31.7173
18.4776 | -35.1516 | -31.7253
18.4837 | -35.1523 | -31.7332
18.4898 | -35.1529 | -31.7411
18.4959 | -35.1535 | -31.749
18.502 | -35.1541 | -31.7568
18.5081 | -35.1548 | -31.7648
18.5142 | -35.1553 | -31.7727
18.5203 | -35.156 | -31.7803
18.5264 | -35.1566 | -31.7877
18.5325 | -35.1572 | -31.7957
18.5386 | -35.1579 | -31.8036
18.5447 | -35.1585 | -31.8112
18.5508 | -35.1591 | -31.8188
18.557 | -35.1597 | -31.8268
18.5631 | -35.1603 | -31.8345
18.5692 | -35.1609 | -31.8422
18.5753 | -35.1616 | -31.8497
18.5814 | -35.1622 | -31.8576
18.5875 | -35.1628 | -31.8656
18.5936 | -35.1634 | -31.8737
18.5997 | -35.164 | -31.8815
18.6058 | -35.1646 | -31.8892
18.6119 | -35.1652 | -31.897
18.618 | -35.1659 | -31.9049
18.6241 | -35.1665 | -31.9128
18.6302 | -35.1672 | -31.9208
18.6363 | -35.1678 | -31.9284
18.6424 | -35.1684 | -31.9359
18.6485 | -35.169 | -31.9437
18.6546 | -35.1696 | -31.9513
18.6607 | -35.1702 | -31.9589
18.6668 | -35.1709 | -31.9666
18.6729 | -35.1715 | -31.9745
18.679 | -35.1721 | -31.9819
18.6851 | -35.1727 | -31.9894
18.6912 | -35.1733 | -31.9972
18.6974 | -35.1739 | -32.0048
18.7035 | -35.1745 | -32.0126
18.7096 | -35.1752 | -32.0203
18.7157 | -35.1758 | -32.0279
18.7218 | -35.1764 | -32.0357
18.7279 | -35.1771 | -32.0434
18.734 | -35.1777 | -32.0511
18.7401 | -35.1783 | -32.0585
18.7462 | -35.1789 | -32.0657
18.7523 | -35.1796 | -32.0735
18.7584 | -35.1802 | -32.0814
18.7645 | -35.1809 | -32.089
18.7706 | -35.1815 | -32.0967
18.7767 | -35.1821 | -32.1044
18.7828 | -35.1827 | -32.1117
18.7889 | -35.1833 | -32.1193
18.795 | -35.1839 | -32.1269
18.8011 | -35.1845 | -32.1347
18.8072 | -35.1852 | -32.1424
18.8133 | -35.1858 | -32.1502
18.8194 | -35.1864 | -32.1577
18.8255 | -35.187 | -32.1652
18.8316 | -35.1876 | -32.173
18.8377 | -35.1882 | -32.1805
18.8439 | -35.1888 | -32.1882
18.85 | -35.1894 | -32.196
18.8561 | -35.19 | -32.2036
18.8622 | -35.1906 | -32.2114
18.8683 | -35.1912 | -32.2188
18.8744 | -35.1919 | -32.2264
18.8805 | -35.1925 | -32.2342
18.8866 | -35.1931 | -32.242
18.8927 | -35.1938 | -32.2493
18.8988 | -35.1944 | -32.257
18.9049 | -35.195 | -32.2647
18.911 | -35.1957 | -32.2722
18.9171 | -35.1963 | -32.2799
18.9232 | -35.197 | -32.2873
18.9293 | -35.1976 | -32.295
18.9354 | -35.1982 | -32.3026
18.9415 | -35.1988 | -32.3101
18.9476 | -35.1994 | -32.3178
18.9537 | -35.2 | -32.3255
18.9598 | -35.2006 | -32.3326
18.9659 | -35.2013 | -32.3402
18.972 | -35.2019 | -32.3478
18.9781 | -35.2025 | -32.3553
18.9843 | -35.2031 | -32.3629
18.9904 | -35.2037 | -32.3706
18.9965 | -35.2043 | -32.378
19.0026 | -35.205 | -32.3858
19.0087 | -35.2056 | -32.3935
19.0148 | -35.2062 | -32.4011
19.0209 | -35.2068 | -32.409
19.027 | -35.2075 | -32.4168
19.0331 | -35.2081 | -32.4243
19.0392 | -35.2087 | -32.4319
19.0453 | -35.2093 | -32.4395
19.0514 | -35.2099 | -32.447
19.0575 | -35.2105 | -32.4547
19.0636 | -35.2111 | -32.4617
19.0697 | -35.2118 | -32.4695
19.0758 | -35.2124 | -32.4771
19.0819 | -35.213 | -32.4849
19.088 | -35.2136 | -32.4922
19.0941 | -35.2142 | -32.4996
19.1002 | -35.2149 | -32.5068
19.1063 | -35.2155 | -32.5144
19.1124 | -35.2161 | -32.522
19.1185 | -35.2168 | -32.5297
19.1246 | -35.2173 | -32.5373
19.1308 | -35.218 | -32.5446
19.1369 | -35.2186 | -32.5519
19.143 | -35.2192 | -32.5595
19.1491 | -35.2198 | -32.567
19.1552 | -35.2204 | -32.5749
19.1613 | -35.2211 | -32.5824
19.1674 | -35.2217 | -32.5901
19.1735 | -35.2223 | -32.5976
19.1796 | -35.2229 | -32.605
19.1857 | -35.2235 | -32.6126
19.1918 | -35.2241 | -32.6203
19.1979 | -35.2247 | -32.6279
19.204 | -35.2254 | -32.6353
19.2101 | -35.226 | -32.6427
19.2162 | -35.2266 | -32.6504
19.2223 | -35.2272 | -32.6578
19.2284 | -35.2279 | -32.6655
19.2345 | -35.2285 | -32.673
19.2406 | -35.2291 | -32.6805
19.2467 | -35.2297 | -32.688
19.2528 | -35.2303 | -32.6956
19.2589 | -35.2309 | -32.7031
19.265 | -35.2315 | -32.7108
19.2712 | -35.2321 | -32.7184
19.2773 | -35.2328 | -32.7256
19.2834 | -35.2334 | -32.733
19.2895 | -35.234 | -32.7406
19.2956 | -35.2347 | -32.7483
19.3017 | -35.2353 | -32.7556
19.3078 | -35.2359 | -32.7631
19.3139 | -35.2365 | -32.7704
19.32 | -35.2372 | -32.7779
19.3261 | -35.2378 | -32.7855
19.3322 | -35.2384 | -32.7929
19.3383 | -35.239 | -32.8003
19.3444 | -35.2397 | -32.8076
19.3505 | -35.2403 | -32.815
19.3566 | -35.2409 | -32.8225
19.3627 | -35.2415 | -32.8301
19.3688 | -35.2422 | -32.8377
19.3749 | -35.2428 | -32.8454
19.381 | -35.2434 | -32.8525
19.3871 | -35.244 | -32.86
19.3932 | -35.2446 | -32.8677
19.3993 | -35.2452 | -32.8751
19.4054 | -35.2459 | -32.8827
19.4115 | -35.2465 | -32.8904
19.4177 | -35.2471 | -32.8976
19.4238 | -35.2477 | -32.905
19.4299 | -35.2484 | -32.9124
19.436 | -35.249 | -32.9203
19.4421 | -35.2496 | -32.9274
19.4482 | -35.2503 | -32.935
19.4543 | -35.2509 | -32.942
19.4604 | -35.2515 | -32.9494
19.4665 | -35.2521 | -32.9567
19.4726 | -35.2528 | -32.964
19.4787 | -35.2534 | -32.9714
19.4848 | -35.254 | -32.9788
19.4909 | -35.2547 | -32.9862
19.497 | -35.2553 | -32.9936
19.5031 | -35.2559 | -33.0009
19.5092 | -35.2565 | -33.0084
19.5153 | -35.2571 | -33.016
19.5214 | -35.2578 | -33.0231
19.5275 | -35.2583 | -33.0304
19.5336 | -35.259 | -33.0377
19.5397 | -35.2596 | -33.0452
19.5458 | -35.2602 | -33.0526
19.5519 | -35.2608 | -33.06
19.5581 | -35.2615 | -33.0671
19.5642 | -35.2621 | -33.0744
19.5703 | -35.2627 | -33.0818
19.5764 | -35.2633 | -33.0893
19.5825 | -35.2639 | -33.0966
19.5886 | -35.2645 | -33.104
19.5947 | -35.2651 | -33.1111
19.6008 | -35.2657 | -33.1186
19.6069 | -35.2664 | -33.1259
19.613 | -35.267 | -33.1334
19.6191 | -35.2676 | -33.1409
19.6252 | -35.2682 | -33.1482
19.6313 | -35.2688 | -33.1552
19.6374 | -35.2694 | -33.1626
19.6435 | -35.27 | -33.1698
19.6496 | -35.2707 | -33.177
19.6557 | -35.2713 | -33.1844
19.6618 | -35.2719 | -33.1915
19.6679 | -35.2725 | -33.1989
19.674 | -35.2731 | -33.2062
19.6801 | -35.2737 | -33.2136
19.6862 | -35.2744 | -33.2208
19.6923 | -35.275 | -33.228
19.6984 | -35.2756 | -33.2349
19.7046 | -35.2763 | -33.2423
19.7107 | -35.2769 | -33.2496
19.7168 | -35.2775 | -33.2571
19.7229 | -35.2781 | -33.2645
19.729 | -35.2787 | -33.2719
19.7351 | -35.2793 | -33.2786
19.7412 | -35.2799 | -33.2861
19.7473 | -35.2805 | -33.2936
19.7534 | -35.2811 | -33.3009
19.7595 | -35.2817 | -33.3085
19.7656 | -35.2823 | -33.316
19.7717 | -35.2829 | -33.3227
19.7778 | -35.2835 | -33.3302
19.7839 | -35.2841 | -33.3374
19.79 | -35.2848 | -33.3447
19.7961 | -35.2853 | -33.3524
19.8022 | -35.2859 | -33.3597
19.8083 | -35.2865 | -33.367
19.8144 | -35.2871 | -33.3743
19.8205 | -35.2877 | -33.3817
19.8266 | -35.2883 | -33.389
19.8327 | -35.289 | -33.3964
19.8388 | -35.2896 | -33.4036
19.845 | -35.2902 | -33.4107
19.8511 | -35.2908 | -33.4181
19.8572 | -35.2914 | -33.4256
19.8633 | -35.292 | -33.4328
19.8694 | -35.2927 | -33.4403
19.8755 | -35.2933 | -33.4474
19.8816 | -35.294 | -33.4549
19.8877 | -35.2946 | -33.4621
19.8938 | -35.2952 | -33.4693
19.8999 | -35.2958 | -33.4765
19.906 | -35.2965 | -33.484
19.9121 | -35.2971 | -33.491
19.9182 | -35.2977 | -33.4981
19.9243 | -35.2984 | -33.5057
19.9304 | -35.299 | -33.5131
19.9365 | -35.2995 | -33.5207
19.9426 | -35.3002 | -33.5277
19.9487 | -35.3008 | -33.5348
19.9548 | -35.3014 | -33.5423
19.9609 | -35.3021 | -33.5495
19.967 | -35.3027 | -33.5567
19.9731 | -35.3033 | -33.564
19.9792 | -35.3039 | -33.5709
19.9853 | -35.3046 | -33.5782
19.9915 | -35.3052 | -33.5856
19.9976 | -35.3058 | -33.5927
20.0037 | -35.3064 | -33.6
20.0098 | -35.307 | -33.6073
20.0159 | -35.3076 | -33.6142
20.022 | -35.3082 | -33.6217
20.0281 | -35.3089 | -33.6291
20.0342 | -35.3095 | -33.6361
20.0403 | -35.3101 | -33.643
20.0464 | -35.3107 | -33.6505
20.0525 | -35.3113 | -33.6575
20.0586 | -35.312 | -33.6646
20.0647 | -35.3126 | -33.6721
20.0708 | -35.3132 | -33.6795
20.0769 | -35.3138 | -33.6867
20.083 | -35.3144 | -33.6935
20.0891 | -35.3149 | -33.7007
20.0952 | -35.3156 | -33.7079
20.1013 | -35.3161 | -33.7152
20.1074 | -35.3167 | -33.7227
20.1135 | -35.3174 | -33.7298
20.1196 | -35.318 | -33.7367
20.1257 | -35.3186 | -33.7441
20.1319 | -35.3192 | -33.7513
20.138 | -35.3198 | -33.7584
20.1441 | -35.3204 | -33.7657
20.1502 | -35.321 | -33.7731
20.1563 | -35.3216 | -33.7799
20.1624 | -35.3222 | -33.787
20.1685 | -35.3229 | -33.7944
20.1746 | -35.3235 | -33.8017
20.1807 | -35.324 | -33.8089
20.1868 | -35.3247 | -33.8161
20.1929 | -35.3253 | -33.8233
20.199 | -35.3259 | -33.8307
20.2051 | -35.3265 | -33.8379
20.2112 | -35.3271 | -33.845
20.2173 | -35.3277 | -33.8523
20.2234 | -35.3283 | -33.8595
20.2295 | -35.3289 | -33.8665
20.2356 | -35.3295 | -33.8738
20.2417 | -35.3301 | -33.8808
20.2478 | -35.3307 | -33.8879
20.2539 | -35.3313 | -33.8953
20.26 | -35.3319 | -33.9023
20.2661 | -35.3325 | -33.9096
20.2723 | -35.3331 | -33.9169
20.2784 | -35.3337 | -33.924
20.2845 | -35.3343 | -33.9311
20.2906 | -35.3349 | -33.9383
20.2967 | -35.3355 | -33.9451
20.3028 | -35.3361 | -33.9522
20.3089 | -35.3367 | -33.9596
20.315 | -35.3374 | -33.9667
20.3211 | -35.3379 | -33.9739
20.3272 | -35.3386 | -33.9809
20.3333 | -35.3392 | -33.9878
20.3394 | -35.3398 | -33.9946
20.3455 | -35.3404 | -34.0016
20.3516 | -35.341 | -34.0086
20.3577 | -35.3416 | -34.0156
20.3638 | -35.3423 | -34.0227
20.3699 | -35.3429 | -34.0296
20.376 | -35.3435 | -34.0367
20.3821 | -35.3442 | -34.0436
20.3882 | -35.3448 | -34.0508
20.3943 | -35.3454 | -34.058
20.4004 | -35.346 | -34.0645
20.4065 | -35.3466 | -34.0715
20.4126 | -35.3472 | -34.0786
20.4188 | -35.3478 | -34.086
20.4249 | -35.3484 | -34.0931
20.431 | -35.349 | -34.1004
20.4371 | -35.3496 | -34.1069
20.4432 | -35.3502 | -34.114
20.4493 | -35.3508 | -34.1211
20.4554 | -35.3514 | -34.1282
20.4615 | -35.352 | -34.1355
20.4676 | -35.3526 | -34.1426
20.4737 | -35.3532 | -34.1494
20.4798 | -35.3538 | -34.1566
20.4859 | -35.3545 | -34.1637
20.492 | -35.3551 | -34.1707
20.4981 | -35.3557 | -34.1777
20.5042 | -35.3563 | -34.1846
20.5103 | -35.3569 | -34.1913
20.5164 | -35.3575 | -34.1984
20.5225 | -35.3581 | -34.2058
20.5286 | -35.3586 | -34.2129
20.5347 | -35.3592 | -34.2198
20.5408 | -35.3598 | -34.2268
20.5469 | -35.3604 | -34.2338
20.553 | -35.3611 | -34.2408
20.5592 | -35.3616 | -34.2477
20.5653 | -35.3622 | -34.2548
20.5714 | -35.3628 | -34.2616
20.5775 | -35.3634 | -34.2684
20.5836 | -35.364 | -34.2756
20.5897 | -35.3646 | -34.2825
20.5958 | -35.3652 | -34.2896
20.6019 | -35.3658 | -34.2966
20.608 | -35.3664 | -34.3035
20.6141 | -35.367 | -34.3102
20.6202 | -35.3676 | -34.3174
20.6263 | -35.3682 | -34.3242
20.6324 | -35.3688 | -34.3312
20.6385 | -35.3693 | -34.338
20.6446 | -35.3699 | -34.3453
20.6507 | -35.3705 | -34.3522
20.6568 | -35.3711 | -34.3592
20.6629 | -35.3717 | -34.3662
20.669 | -35.3723 | -34.3731
20.6751 | -35.3729 | -34.3801
20.6812 | -35.3735 | -34.3866
20.6873 | -35.3741 | -34.3937
20.6934 | -35.3748 | -34.4007
20.6995 | -35.3754 | -34.4076
20.7057 | -35.376 | -34.4147
20.7118 | -35.3766 | -34.4216
20.7179 | -35.3772 | -34.428
20.724 | -35.3777 | -34.4346
20.7301 | -35.3783 | -34.4417
20.7362 | -35.3789 | -34.4488
20.7423 | -35.3795 | -34.4557
20.7484 | -35.3801 | -34.463
20.7545 | -35.3807 | -34.47
20.7606 | -35.3813 | -34.477
20.7667 | -35.3819 | -34.4839
20.7728 | -35.3825 | -34.4909
20.7789 | -35.3831 | -34.4979
20.785 | -35.3837 | -34.5049
20.7911 | -35.3843 | -34.5118
20.7972 | -35.3849 | -34.5186
20.8033 | -35.3855 | -34.5254
20.8094 | -35.3861 | -34.5324
20.8155 | -35.3867 | -34.5395
20.8216 | -35.3872 | -34.5465
20.8277 | -35.3878 | -34.5532
20.8338 | -35.3884 | -34.5596
20.8399 | -35.3891 | -34.5668
20.8461 | -35.3897 | -34.5733
20.8522 | -35.3903 | -34.58
20.8583 | -35.3909 | -34.5864
20.8644 | -35.3915 | -34.5932
20.8705 | -35.3921 | -34.6004
20.8766 | -35.3927 | -34.6075
20.8827 | -35.3934 | -34.6144
20.8888 | -35.3939 | -34.6214
20.8949 | -35.3945 | -34.628
20.901 | -35.3951 | -34.6347
20.9071 | -35.3957 | -34.6415
20.9132 | -35.3963 | -34.6485
20.9193 | -35.3969 | -34.6555
20.9254 | -35.3975 | -34.6624
20.9315 | -35.3981 | -34.6688
20.9376 | -35.3987 | -34.6758
20.9437 | -35.3993 | -34.6828
20.9498 | -35.3999 | -34.6895
20.9559 | -35.4006 | -34.6961
20.962 | -35.4012 | -34.7032
20.9681 | -35.4018 | -34.71
20.9742 | -35.4024 | -34.7167
20.9803 | -35.403 | -34.7234
20.9864 | -35.4036 | -34.7304
20.9926 | -35.4042 | -34.7373
20.9987 | -35.4048 | -34.7438
21.0048 | -35.4054 | -34.7506
21.0109 | -35.406 | -34.7575
21.017 | -35.4066 | -34.7644
21.0231 | -35.4072 | -34.7713
21.0292 | -35.4078 | -34.7782
21.0353 | -35.4084 | -34.7848
21.0414 | -35.409 | -34.7915
21.0475 | -35.4096 | -34.7985
21.0536 | -35.4102 | -34.8052
21.0597 | -35.4108 | -34.8118
21.0658 | -35.4114 | -34.8187
21.0719 | -35.4121 | -34.8259
21.078 | -35.4128 | -34.8327
21.0841 | -35.4133 | -34.8395
21.0902 | -35.414 | -34.8468
21.0963 | -35.4145 | -34.8532
21.1024 | -35.4151 | -34.86
21.1085 | -35.4156 | -34.8665
21.1146 | -35.4162 | -34.8736
21.1207 | -35.4168 | -34.8804
21.1268 | -35.4174 | -34.8871
21.133 | -35.418 | -34.8938
21.1391 | -35.4185 | -34.9004
21.1452 | -35.4191 | -34.9074
21.1513 | -35.4198 | -34.914
21.1574 | -35.4204 | -34.921
21.1635 | -35.4209 | -34.9275
21.1696 | -35.4216 | -34.9341
21.1757 | -35.4222 | -34.9405
21.1818 | -35.4228 | -34.9475
21.1879 | -35.4234 | -34.9544
21.194 | -35.424 | -34.9612
21.2001 | -35.4246 | -34.9677
21.2062 | -35.4251 | -34.9745
21.2123 | -35.4257 | -34.981
21.2184 | -35.4263 | -34.988
21.2245 | -35.4269 | -34.9946
21.2306 | -35.4275 | -35.001
21.2367 | -35.4281 | -35.0079
21.2428 | -35.4288 | -35.0145
21.2489 | -35.4293 | -35.0209
21.255 | -35.4299 | -35.0277
21.2611 | -35.4305 | -35.0345
21.2672 | -35.4311 | -35.0415
21.2733 | -35.4317 | -35.0482
21.2795 | -35.4323 | -35.0548
21.2856 | -35.4329 | -35.0617
21.2917 | -35.4335 | -35.0685
21.2978 | -35.4341 | -35.0753
21.3039 | -35.4347 | -35.0819
21.31 | -35.4353 | -35.0885
21.3161 | -35.4359 | -35.0947
21.3222 | -35.4365 | -35.1013
21.3283 | -35.4371 | -35.1082
21.3344 | -35.4376 | -35.1147
21.3405 | -35.4382 | -35.1215
21.3466 | -35.4389 | -35.1283
21.3527 | -35.4394 | -35.1345
21.3588 | -35.44 | -35.141
21.3649 | -35.4406 | -35.1476
21.371 | -35.4412 | -35.1542
21.3771 | -35.4418 | -35.1607
21.3832 | -35.4424 | -35.1673
21.3893 | -35.443 | -35.1736
21.3954 | -35.4436 | -35.1803
21.4015 | -35.4442 | -35.187
21.4076 | -35.4448 | -35.1937
21.4137 | -35.4454 | -35.2005
21.4199 | -35.4459 | -35.2071
21.426 | -35.4465 | -35.2138
21.4321 | -35.4471 | -35.2205
21.4382 | -35.4477 | -35.227
21.4443 | -35.4483 | -35.2336
21.4504 | -35.4489 | -35.2404
21.4565 | -35.4494 | -35.247
21.4626 | -35.4501 | -35.2536
21.4687 | -35.4507 | -35.2602
21.4748 | -35.4513 | -35.2668
21.4809 | -35.4518 | -35.2734
21.487 | -35.4524 | -35.28
21.4931 | -35.453 | -35.2865
21.4992 | -35.4536 | -35.2929
21.5053 | -35.4541 | -35.2998
21.5114 | -35.4547 | -35.3064
21.5175 | -35.4552 | -35.3132
21.5236 | -35.4558 | -35.3198
21.5297 | -35.4563 | -35.3258
21.5358 | -35.4569 | -35.3324
21.5419 | -35.4574 | -35.3393
21.548 | -35.458 | -35.3461
21.5541 | -35.4586 | -35.3528
21.5602 | -35.4591 | -35.3595
21.5664 | -35.4597 | -35.366
21.5725 | -35.4603 | -35.3725
21.5786 | -35.4608 | -35.3791
21.5847 | -35.4614 | -35.3858
21.5908 | -35.462 | -35.3924
21.5969 | -35.4626 | -35.3988
21.603 | -35.4632 | -35.4055
21.6091 | -35.4638 | -35.4124
21.6152 | -35.4644 | -35.4192
21.6213 | -35.465 | -35.426
21.6274 | -35.4656 | -35.4327
21.6335 | -35.4661 | -35.4389
21.6396 | -35.4667 | -35.4455
21.6457 | -35.4673 | -35.4523
21.6518 | -35.4678 | -35.4589
21.6579 | -35.4684 | -35.4651
21.664 | -35.469 | -35.4719
21.6701 | -35.4696 | -35.4783
21.6762 | -35.4702 | -35.4852
21.6823 | -35.4708 | -35.4918
21.6884 | -35.4714 | -35.4985
21.6945 | -35.472 | -35.505
21.7006 | -35.4725 | -35.5118
21.7068 | -35.4731 | -35.5181
21.7129 | -35.4737 | -35.5247
21.719 | -35.4743 | -35.5314
21.7251 | -35.4749 | -35.538
21.7312 | -35.4755 | -35.5445
21.7373 | -35.4761 | -35.5506
21.7434 | -35.4767 | -35.5573
21.7495 | -35.4772 | -35.5637
21.7556 | -35.4778 | -35.5703
21.7617 | -35.4784 | -35.5768
21.7678 | -35.479 | -35.5836
21.7739 | -35.4796 | -35.5897
21.78 | -35.4802 | -35.5961
21.7861 | -35.4808 | -35.6028
21.7922 | -35.4813 | -35.6094
21.7983 | -35.4819 | -35.6159
21.8044 | -35.4826 | -35.6227
21.8105 | -35.4831 | -35.629
21.8166 | -35.4837 | -35.6356
21.8227 | -35.4843 | -35.6421
21.8288 | -35.4849 | -35.6486
21.8349 | -35.4855 | -35.6552
21.841 | -35.4861 | -35.6618
21.8471 | -35.4866 | -35.6681
21.8533 | -35.4872 | -35.6744
21.8594 | -35.4878 | -35.6812
21.8655 | -35.4884 | -35.6879
21.8716 | -35.489 | -35.6945
21.8777 | -35.4895 | -35.7008
21.8838 | -35.4901 | -35.7071
21.8899 | -35.4907 | -35.7135
21.896 | -35.4913 | -35.7201
21.9021 | -35.4919 | -35.7267
21.9082 | -35.4924 | -35.7331
21.9143 | -35.493 | -35.7391
21.9204 | -35.4936 | -35.7457
21.9265 | -35.4942 | -35.7521
21.9326 | -35.4947 | -35.7586
21.9387 | -35.4953 | -35.7651
21.9448 | -35.4959 | -35.7719
21.9509 | -35.4965 | -35.7779
21.957 | -35.4971 | -35.7844
21.9631 | -35.4977 | -35.7912
21.9692 | -35.4982 | -35.7978
21.9753 | -35.4988 | -35.8042
21.9814 | -35.4994 | -35.8105
21.9875 | -35.5 | -35.8168
21.9937 | -35.5005 | -35.8232
21.9998 | -35.5011 | -35.8298
22.0059 | -35.5017 | -35.8364
22.012 | -35.5023 | -35.8428
22.0181 | -35.5029 | -35.8491
22.0242 | -35.5035 | -35.8556
22.0303 | -35.5041 | -35.8621
22.0364 | -35.5046 | -35.8685
22.0425 | -35.5052 | -35.8751
22.0486 | -35.5058 | -35.8814
22.0547 | -35.5064 | -35.8875
22.0608 | -35.507 | -35.8938
22.0669 | -35.5076 | -35.9002
22.073 | -35.5081 | -35.9065
22.0791 | -35.5087 | -35.9129
22.0852 | -35.5093 | -35.9193
22.0913 | -35.5099 | -35.9255
22.0974 | -35.5105 | -35.9318
22.1035 | -35.511 | -35.9382
22.1096 | -35.5116 | -35.9449
22.1157 | -35.5122 | -35.9513
22.1218 | -35.5128 | -35.958
22.1279 | -35.5133 | -35.964
22.134 | -35.5139 | -35.9703
22.1402 | -35.5145 | -35.9767
22.1463 | -35.5151 | -35.9831
22.1524 | -35.5156 | -35.9896
22.1585 | -35.5162 | -35.9961
22.1646 | -35.5168 | -36.0021
22.1707 | -35.5174 | -36.009
22.1768 | -35.518 | -36.0155
22.1829 | -35.5186 | -36.022
22.189 | -35.5191 | -36.0284
22.1951 | -35.5197 | -36.0346
22.2012 | -35.5203 | -36.0412
22.2073 | -35.5208 | -36.0476
22.2134 | -35.5215 | -36.0541
22.2195 | -35.5221 | -36.0605
22.2256 | -35.5226 | -36.067
22.2317 | -35.5232 | -36.0729
22.2378 | -35.5238 | -36.0794
22.2439 | -35.5244 | -36.0858
22.25 | -35.525 | -36.0923
22.2561 | -35.5256 | -36.0988
22.2622 | -35.5261 | -36.1053
22.2683 | -35.5267 | -36.1112
22.2744 | -35.5273 | -36.1176
22.2806 | -35.5279 | -36.124
22.2867 | -35.5284 | -36.1302
22.2928 | -35.529 | -36.1365
22.2989 | -35.5296 | -36.143
22.305 | -35.5301 | -36.1493
22.3111 | -35.5307 | -36.1557
22.3172 | -35.5313 | -36.1621
22.3233 | -35.5319 | -36.1685
22.3294 | -35.5325 | -36.1748
22.3355 | -35.5331 | -36.1809
22.3416 | -35.5336 | -36.1873
22.3477 | -35.5342 | -36.1939
22.3538 | -35.5348 | -36.2001
22.3599 | -35.5353 | -36.2064
22.366 | -35.5359 | -36.2128
22.3721 | -35.5365 | -36.2189
22.3782 | -35.5371 | -36.2254
22.3843 | -35.5377 | -36.232
22.3904 | -35.5382 | -36.2386
22.3965 | -35.5388 | -36.2451
22.4026 | -35.5394 | -36.2513
22.4087 | -35.5399 | -36.2576
22.4148 | -35.5405 | -36.264
22.4209 | -35.541 | -36.2705
22.4271 | -35.5416 | -36.2768
22.4332 | -35.5422 | -36.2829
22.4393 | -35.5428 | -36.2892
22.4454 | -35.5433 | -36.295
22.4515 | -35.5438 | -36.3014
22.4576 | -35.5445 | -36.3078
22.4637 | -35.545 | -36.3141
22.4698 | -35.5456 | -36.3204
22.4759 | -35.5462 | -36.3264
22.482 | -35.5468 | -36.3329
22.4881 | -35.5473 | -36.3392
22.4942 | -35.5479 | -36.3455
22.5003 | -35.5485 | -36.3518
22.5064 | -35.549 | -36.3582
22.5125 | -35.5495 | -36.3645
22.5186 | -35.5501 | -36.3708
22.5247 | -35.5507 | -36.3773
22.5308 | -35.5512 | -36.3838
22.5369 | -35.5518 | -36.3904
22.543 | -35.5524 | -36.3967
22.5491 | -35.553 | -36.4027
22.5552 | -35.5535 | -36.4092
22.5613 | -35.5541 | -36.4156
22.5675 | -35.5546 | -36.422
22.5736 | -35.5552 | -36.4285
22.5797 | -35.5557 | -36.4347
22.5858 | -35.5563 | -36.4407
22.5919 | -35.5569 | -36.447
22.598 | -35.5574 | -36.4535
22.6041 | -35.558 | -36.4597
22.6102 | -35.5586 | -36.466
22.6163 | -35.5591 | -36.4722
22.6224 | -35.5597 | -36.4785
22.6285 | -35.5603 | -36.4849
22.6346 | -35.5608 | -36.4911
22.6407 | -35.5614 | -36.4975
22.6468 | -35.562 | -36.5038
22.6529 | -35.5625 | -36.5096
22.659 | -35.5631 | -36.5159
22.6651 | -35.5636 | -36.5222
22.6712 | -35.5642 | -36.5284
22.6773 | -35.5648 | -36.5347
22.6834 | -35.5654 | -36.5411
22.6895 | -35.5659 | -36.5472
22.6956 | -35.5665 | -36.5535
22.7017 | -35.5671 | -36.5598
22.7079 | -35.5676 | -36.5662
22.714 | -35.5682 | -36.5724
22.7201 | -35.5688 | -36.5788
22.7262 | -35.5693 | -36.5852
22.7323 | -35.5699 | -36.5915
22.7384 | -35.5704 | -36.5978
22.7445 | -35.571 | -36.6043
22.7506 | -35.5716 | -36.6105
22.7567 | -35.5721 | -36.617
22.7628 | -35.5727 | -36.6229
22.7689 | -35.5733 | -36.6291
22.775 | -35.5738 | -36.6355
22.7811 | -35.5744 | -36.6417
22.7872 | -35.575 | -36.648
22.7933 | -35.5755 | -36.654
22.7994 | -35.5761 | -36.6601
22.8055 | -35.5767 | -36.6665
22.8116 | -35.5772 | -36.6727
22.8177 | -35.5778 | -36.679
22.8238 | -35.5783 | -36.6854
22.8299 | -35.5789 | -36.6913
22.836 | -35.5795 | -36.6978
22.8421 | -35.5801 | -36.704
22.8482 | -35.5806 | -36.7103
22.8544 | -35.5812 | -36.7166
22.8605 | -35.5818 | -36.7228
22.8666 | -35.5823 | -36.7291
22.8727 | -35.5829 | -36.7353
22.8788 | -35.5835 | -36.7415
22.8849 | -35.584 | -36.7477
22.891 | -35.5846 | -36.7541
22.8971 | -35.5852 | -36.7602
22.9032 | -35.5857 | -36.7663
22.9093 | -35.5863 | -36.7726
22.9154 | -35.5868 | -36.7789
22.9215 | -35.5873 | -36.7851
22.9276 | -35.5879 | -36.7912
22.9337 | -35.5884 | -36.7972
22.9398 | -35.589 | -36.8034
22.9459 | -35.5896 | -36.8097
22.952 | -35.5901 | -36.8158
22.9581 | -35.5907 | -36.8221
22.9642 | -35.5913 | -36.8286
22.9703 | -35.5918 | -36.8348
22.9764 | -35.5924 | -36.8411
22.9825 | -35.593 | -36.8473
22.9886 | -35.5935 | -36.8534
22.9948 | -35.5941 | -36.8596
23.0009 | -35.5946 | -36.8658
23.007 | -35.5952 | -36.872
23.0131 | -35.5957 | -36.8783
23.0192 | -35.5963 | -36.8843
23.0253 | -35.5969 | -36.8906
23.0314 | -35.5974 | -36.8967
23.0375 | -35.598 | -36.9029
23.0436 | -35.5986 | -36.9089
23.0497 | -35.5991 | -36.9153
23.0558 | -35.5997 | -36.9214
23.0619 | -35.6003 | -36.9278
23.068 | -35.6009 | -36.9341
23.0741 | -35.6014 | -36.9399
23.0802 | -35.602 | -36.946
23.0863 | -35.6025 | -36.9521
23.0924 | -35.6031 | -36.9583
23.0985 | -35.6036 | -36.9646
23.1046 | -35.6042 | -36.9708
23.1107 | -35.6047 | -36.9769
23.1168 | -35.6053 | -36.9833
23.1229 | -35.6058 | -36.9896
23.129 | -35.6064 | -36.9961
23.1351 | -35.607 | -37.0023
23.1413 | -35.6075 | -37.0085
23.1474 | -35.6081 | -37.0144
23.1535 | -35.6086 | -37.0207
23.1596 | -35.6092 | -37.0267
23.1657 | -35.6097 | -37.0329
23.1718 | -35.6103 | -37.0391
23.1779 | -35.6108 | -37.0455
23.184 | -35.6114 | -37.0513
23.1901 | -35.6119 | -37.0576
23.1962 | -35.6125 | -37.0636
23.2023 | -35.613 | -37.0696
23.2084 | -35.6136 | -37.076
23.2145 | -35.6142 | -37.0822
23.2206 | -35.6147 | -37.0884
23.2267 | -35.6153 | -37.0946
23.2328 | -35.6158 | -37.1008
23.2389 | -35.6164 | -37.1071
23.245 | -35.6169 | -37.1135
23.2511 | -35.6175 | -37.1197
23.2572 | -35.6181 | -37.1258
23.2633 | -35.6187 | -37.1319
23.2694 | -35.6192 | -37.1382
23.2755 | -35.6198 | -37.1449
23.2817 | -35.6204 | -37.1511
23.2878 | -35.6209 | -37.1569
23.2939 | -35.6214 | -37.1631
23.3 | -35.622 | -37.1695
23.3061 | -35.6225 | -37.1757
23.3122 | -35.6231 | -37.1819
23.3183 | -35.6237 | -37.1881
23.3244 | -35.6242 | -37.1941
23.3305 | -35.6248 | -37.2004
23.3366 | -35.6253 | -37.2067
23.3427 | -35.6259 | -37.2128
23.3488 | -35.6264 | -37.2191
23.3549 | -35.627 | -37.2253
23.361 | -35.6276 | -37.2315
23.3671 | -35.6282 | -37.2377
23.3732 | -35.6288 | -37.2438
23.3793 | -35.6293 | -37.2499
23.3854 | -35.6299 | -37.2563
23.3915 | -35.6304 | -37.2623
23.3976 | -35.631 | -37.2686
23.4037 | -35.6316 | -37.2745
23.4098 | -35.6321 | -37.2806
23.4159 | -35.6327 | -37.2866
23.422 | -35.6333 | -37.2927
23.4282 | -35.6338 | -37.2987
23.4343 | -35.6344 | -37.3048
23.4404 | -35.6349 | -37.3112
23.4465 | -35.6354 | -37.3175
23.4526 | -35.636 | -37.3238
23.4587 | -35.6366 | -37.3298
23.4648 | -35.6371 | -37.3357
23.4709 | -35.6376 | -37.3419
23.477 | -35.6381 | -37.3482
23.4831 | -35.6387 | -37.3543
23.4892 | -35.6392 | -37.3606
23.4953 | -35.6398 | -37.3669
23.5014 | -35.6403 | -37.373
23.5075 | -35.6409 | -37.3791
23.5136 | -35.6415 | -37.3851
23.5197 | -35.642 | -37.3913
23.5258 | -35.6425 | -37.3972
23.5319 | -35.6431 | -37.4033
23.538 | -35.6436 | -37.4094
23.5441 | -35.6442 | -37.4155
23.5502 | -35.6447 | -37.422
23.5563 | -35.6453 | -37.428
23.5624 | -35.6458 | -37.4343
23.5686 | -35.6463 | -37.4403
23.5747 | -35.6469 | -37.4463
23.5808 | -35.6475 | -37.4524
23.5869 | -35.648 | -37.4584
23.593 | -35.6486 | -37.4644
23.5991 | -35.6492 | -37.4706
23.6052 | -35.6497 | -37.4764
23.6113 | -35.6502 | -37.4825
23.6174 | -35.6507 | -37.4887
23.6235 | -35.6513 | -37.4948
23.6296 | -35.6518 | -37.501
23.6357 | -35.6524 | -37.5074
23.6418 | -35.6529 | -37.5133
23.6479 | -35.6535 | -37.5194
23.654 | -35.654 | -37.5255
23.6601 | -35.6546 | -37.5315
23.6662 | -35.6551 | -37.5377
23.6723 | -35.6557 | -37.5439
23.6784 | -35.6562 | -37.5499
23.6845 | -35.6568 | -37.5559
23.6906 | -35.6573 | -37.5622
23.6967 | -35.6579 | -37.5683
23.7028 | -35.6585 | -37.5744
23.7089 | -35.659 | -37.5807
23.7151 | -35.6596 | -37.5869
23.7212 | -35.6601 | -37.5929
23.7273 | -35.6607 | -37.5992
23.7334 | -35.6612 | -37.6052
23.7395 | -35.6618 | -37.6113
23.7456 | -35.6623 | -37.6174
23.7517 | -35.6628 | -37.6235
23.7578 | -35.6634 | -37.6296
23.7639 | -35.6639 | -37.6356
23.77 | -35.6645 | -37.6418
23.7761 | -35.665 | -37.6478
23.7822 | -35.6656 | -37.654
23.7883 | -35.6661 | -37.66
23.7944 | -35.6666 | -37.6662
23.8005 | -35.6672 | -37.6723
23.8066 | -35.6677 | -37.6784
23.8127 | -35.6683 | -37.6845
23.8188 | -35.6688 | -37.6907
23.8249 | -35.6694 | -37.6969
23.831 | -35.6699 | -37.7031
23.8371 | -35.6705 | -37.7091
23.8432 | -35.671 | -37.7152
23.8493 | -35.6715 | -37.7211
23.8555 | -35.6721 | -37.7273
23.8616 | -35.6726 | -37.7334
23.8677 | -35.6732 | -37.7395
23.8738 | -35.6737 | -37.7457
23.8799 | -35.6742 | -37.7519
23.886 | -35.6747 | -37.7577
23.8921 | -35.6752 | -37.7638
23.8982 | -35.6758 | -37.7699
23.9043 | -35.6763 | -37.7763
23.9104 | -35.6769 | -37.7825
23.9165 | -35.6774 | -37.7884
23.9226 | -35.6779 | -37.7943
23.9287 | -35.6785 | -37.8004
23.9348 | -35.679 | -37.8064
23.9409 | -35.6795 | -37.8124
23.947 | -35.68 | -37.8185
23.9531 | -35.6806 | -37.8245
23.9592 | -35.6811 | -37.8306
23.9653 | -35.6816 | -37.8368
23.9714 | -35.6821 | -37.8427
23.9775 | -35.6827 | -37.8489
23.9836 | -35.6832 | -37.855
23.9897 | -35.6838 | -37.861
23.9958 | -35.6843 | -37.8674
24.002 | -35.6848 | -37.8733
24.0081 | -35.6854 | -37.8792
24.0142 | -35.6859 | -37.8852
24.0203 | -35.6865 | -37.8911
24.0264 | -35.687 | -37.8972
24.0325 | -35.6875 | -37.9032
24.0386 | -35.688 | -37.9095
24.0447 | -35.6886 | -37.9154
24.0508 | -35.6891 | -37.9215
24.0569 | -35.6896 | -37.9277
24.063 | -35.6902 | -37.9336
24.0691 | -35.6907 | -37.9395
24.0752 | -35.6913 | -37.9457
24.0813 | -35.6918 | -37.9517
24.0874 | -35.6923 | -37.9578
24.0935 | -35.6928 | -37.964
24.0996 | -35.6933 | -37.9699
24.1057 | -35.6939 | -37.976
24.1118 | -35.6944 | -37.982
24.1179 | -35.6949 | -37.9882
24.124 | -35.6954 | -37.9941
24.1301 | -35.696 | -38
24.1362 | -35.6965 | -38.0059
24.1424 | -35.697 | -38.012
24.1485 | -35.6975 | -38.0179
24.1546 | -35.698 | -38.0241
24.1607 | -35.6986 | -38.0301
24.1668 | -35.6991 | -38.036
24.1729 | -35.6996 | -38.0421
24.179 | -35.7001 | -38.0481
24.1851 | -35.7007 | -38.0544
24.1912 | -35.7012 | -38.0604
24.1973 | -35.7017 | -38.0666
24.2034 | -35.7023 | -38.0724
24.2095 | -35.7028 | -38.0786
24.2156 | -35.7034 | -38.0847
24.2217 | -35.7039 | -38.0907
24.2278 | -35.7045 | -38.0966
24.2339 | -35.705 | -38.1026
24.24 | -35.7055 | -38.1085
24.2461 | -35.706 | -38.1145
24.2522 | -35.7066 | -38.1205
24.2583 | -35.7071 | -38.1266
24.2644 | -35.7077 | -38.1327
24.2705 | -35.7082 | -38.1387
24.2766 | -35.7088 | -38.1447
24.2827 | -35.7093 | -38.1506
24.2889 | -35.7098 | -38.1566
24.295 | -35.7103 | -38.1628
24.3011 | -35.7108 | -38.1689
24.3072 | -35.7114 | -38.1747
24.3133 | -35.7119 | -38.1806
24.3194 | -35.7124 | -38.1869
24.3255 | -35.713 | -38.193
24.3316 | -35.7135 | -38.1991
24.3377 | -35.714 | -38.2051
24.3438 | -35.7144 | -38.2112
24.3499 | -35.7149 | -38.2173
24.356 | -35.7154 | -38.2234
24.3621 | -35.716 | -38.2295
24.3682 | -35.7164 | -38.2353
24.3743 | -35.717 | -38.2416
24.3804 | -35.7174 | -38.2473
24.3865 | -35.718 | -38.2534
24.3926 | -35.7185 | -38.2596
24.3987 | -35.719 | -38.2657
24.4048 | -35.7195 | -38.2716
24.4109 | -35.72 | -38.2775
24.417 | -35.7205 | -38.2835
24.4231 | -35.7211 | -38.2897
24.4293 | -35.7216 | -38.2956
24.4354 | -35.7221 | -38.3014
24.4415 | -35.7226 | -38.3071
24.4476 | -35.7232 | -38.3127
24.4537 | -35.7237 | -38.3186
24.4598 | -35.7243 | -38.3248
24.4659 | -35.7248 | -38.331
24.472 | -35.7253 | -38.337
24.4781 | -35.7258 | -38.3431
24.4842 | -35.7263 | -38.349
24.4903 | -35.7268 | -38.3549
24.4964 | -35.7273 | -38.361
24.5025 | -35.7279 | -38.3671
24.5086 | -35.7284 | -38.3731
24.5147 | -35.7289 | -38.3788
24.5208 | -35.7295 | -38.3842
24.5269 | -35.73 | -38.3899
24.533 | -35.7305 | -38.3957
24.5391 | -35.731 | -38.4017
24.5452 | -35.7316 | -38.4077
24.5513 | -35.732 | -38.4134
24.5574 | -35.7326 | -38.4195
24.5635 | -35.7331 | -38.4252
24.5696 | -35.7336 | -38.4312
24.5758 | -35.7342 | -38.4372
24.5819 | -35.7347 | -38.4433
24.588 | -35.7352 | -38.4493
24.5941 | -35.7357 | -38.4551
24.6002 | -35.7362 | -38.4611
24.6063 | -35.7368 | -38.4672
24.6124 | -35.7373 | -38.4733
24.6185 | -35.7378 | -38.4792
24.6246 | -35.7382 | -38.4851
24.6307 | -35.7388 | -38.4912
24.6368 | -35.7393 | -38.4972
24.6429 | -35.7398 | -38.5032
24.649 | -35.7403 | -38.509
24.6551 | -35.7408 | -38.515
24.6612 | -35.7413 | -38.5206
24.6673 | -35.7418 | -38.5267
24.6734 | -35.7423 | -38.5324
24.6795 | -35.7428 | -38.5384
24.6856 | -35.7434 | -38.5446
24.6917 | -35.7439 | -38.5506
24.6978 | -35.7444 | -38.5562
24.7039 | -35.7449 | -38.5624
24.71 | -35.7454 | -38.5685
24.7162 | -35.7459 | -38.5745
24.7223 | -35.7464 | -38.5806
24.7284 | -35.7468 | -38.5865
24.7345 | -35.7474 | -38.5925
24.7406 | -35.7479 | -38.5987
24.7467 | -35.7484 | -38.6045
24.7528 | -35.7489 | -38.6105
24.7589 | -35.7494 | -38.6162
24.765 | -35.7499 | -38.6219
24.7711 | -35.7504 | -38.6276
24.7772 | -35.7509 | -38.6335
24.7833 | -35.7514 | -38.6394
24.7894 | -35.7519 | -38.6453
24.7955 | -35.7524 | -38.6513
24.8016 | -35.7529 | -38.6572
24.8077 | -35.7534 | -38.6633
24.8138 | -35.7539 | -38.6692
24.8199 | -35.7544 | -38.6753
24.826 | -35.7549 | -38.6812
24.8321 | -35.7554 | -38.6874
24.8382 | -35.7559 | -38.6932
24.8443 | -35.7564 | -38.699
24.8504 | -35.7569 | -38.7049
24.8565 | -35.7574 | -38.711
24.8627 | -35.7579 | -38.717
24.8688 | -35.7584 | -38.723
24.8749 | -35.759 | -38.7287
24.881 | -35.7594 | -38.7347
24.8871 | -35.76 | -38.7409
24.8932 | -35.7605 | -38.7468
24.8993 | -35.761 | -38.7525
24.9054 | -35.7615 | -38.7583
24.9115 | -35.7621 | -38.7643
24.9176 | -35.7625 | -38.7702
24.9237 | -35.7631 | -38.7761
24.9298 | -35.7636 | -38.7819
24.9359 | -35.7641 | -38.7878
24.942 | -35.7645 | -38.7936
24.9481 | -35.7651 | -38.7996
24.9542 | -35.7656 | -38.8056
24.9603 | -35.7661 | -38.8116
24.9664 | -35.7666 | -38.8175
24.9725 | -35.7671 | -38.8236
24.9786 | -35.7676 | -38.8295
24.9847 | -35.768 | -38.8355
24.9908 | -35.7685 | -38.8414
24.9969 | -35.769 | -38.8475
25.0031 | -35.7695 | -38.8534
25.0092 | -35.77 | -38.8591
25.0153 | -35.7705 | -38.8648
25.0214 | -35.771 | -38.871
25.0275 | -35.7715 | -38.8769
25.0336 | -35.7721 | -38.8828
25.0397 | -35.7726 | -38.8888
25.0458 | -35.773 | -38.8945
25.0519 | -35.7735 | -38.9005
25.058 | -35.774 | -38.9064
25.0641 | -35.7745 | -38.9122
25.0702 | -35.775 | -38.918
25.0763 | -35.7755 | -38.9237
25.0824 | -35.7759 | -38.9295
25.0885 | -35.7764 | -38.9353
25.0946 | -35.7769 | -38.9413
25.1007 | -35.7774 | -38.9472
25.1068 | -35.7779 | -38.953
25.1129 | -35.7784 | -38.9588
25.119 | -35.7789 | -38.9646
25.1251 | -35.7794 | -38.9704
25.1312 | -35.7799 | -38.9764
25.1373 | -35.7804 | -38.9826
25.1435 | -35.7809 | -38.9886
25.1496 | -35.7814 | -38.9944
25.1557 | -35.7819 | -39.0004
25.1618 | -35.7824 | -39.0064
25.1679 | -35.7829 | -39.0123
25.174 | -35.7834 | -39.0181
25.1801 | -35.7839 | -39.0241
25.1862 | -35.7844 | -39.0299
25.1923 | -35.7849 | -39.036
25.1984 | -35.7854 | -39.0419
25.2045 | -35.7859 | -39.0477
25.2106 | -35.7864 | -39.0536
25.2167 | -35.7869 | -39.0594
25.2228 | -35.7874 | -39.0652
25.2289 | -35.7879 | -39.071
25.235 | -35.7884 | -39.077
25.2411 | -35.7889 | -39.0829
25.2472 | -35.7893 | -39.0888
25.2533 | -35.7898 | -39.0947
25.2594 | -35.7903 | -39.1005
25.2655 | -35.7907 | -39.1062
25.2716 | -35.7912 | -39.1121
25.2777 | -35.7917 | -39.118
25.2838 | -35.7921 | -39.1241
25.29 | -35.7926 | -39.1299
25.2961 | -35.7931 | -39.1357
25.3022 | -35.7935 | -39.1414
25.3083 | -35.794 | -39.1472
25.3144 | -35.7945 | -39.1531
25.3205 | -35.795 | -39.1588
25.3266 | -35.7954 | -39.1647
25.3327 | -35.7959 | -39.1706
25.3388 | -35.7964 | -39.1764
25.3449 | -35.7969 | -39.1824
25.351 | -35.7974 | -39.1882
25.3571 | -35.7979 | -39.1943
25.3632 | -35.7984 | -39.1999
25.3693 | -35.7988 | -39.2057
25.3754 | -35.7993 | -39.2117
25.3815 | -35.7998 | -39.2175
25.3876 | -35.8003 | -39.2234
25.3937 | -35.8008 | -39.2291
25.3998 | -35.8013 | -39.235
25.4059 | -35.8017 | -39.241
25.412 | -35.8022 | -39.2467
25.4181 | -35.8027 | -39.2528
25.4242 | -35.8032 | -39.2585
25.4304 | -35.8037 | -39.2643
25.4365 | -35.8041 | -39.27
25.4426 | -35.8046 | -39.276
25.4487 | -35.8051 | -39.2821
25.4548 | -35.8056 | -39.2878
25.4609 | -35.8061 | -39.2937
25.467 | -35.8066 | -39.2994
25.4731 | -35.8071 | -39.3054
25.4792 | -35.8076 | -39.3111
25.4853 | -35.8081 | -39.3169
25.4914 | -35.8086 | -39.3228
25.4975 | -35.8091 | -39.3284
25.5036 | -35.8095 | -39.3341
25.5097 | -35.81 | -39.3398
25.5158 | -35.8105 | -39.3457
25.5219 | -35.811 | -39.3515
25.528 | -35.8115 | -39.3573
25.5341 | -35.812 | -39.3631
25.5402 | -35.8124 | -39.3689
25.5463 | -35.8129 | -39.3749
25.5524 | -35.8134 | -39.3808
25.5585 | -35.8139 | -39.3864
25.5646 | -35.8143 | -39.3924
25.5707 | -35.8148 | -39.3983
25.5769 | -35.8153 | -39.404
25.583 | -35.8158 | -39.4101
25.5891 | -35.8162 | -39.416
25.5952 | -35.8167 | -39.4218
25.6013 | -35.8172 | -39.4277
25.6074 | -35.8177 | -39.4335
25.6135 | -35.8182 | -39.4393
25.6196 | -35.8187 | -39.4453
25.6257 | -35.8191 | -39.4511
25.6318 | -35.8196 | -39.4568
25.6379 | -35.8201 | -39.4627
25.644 | -35.8206 | -39.4683
25.6501 | -35.821 | -39.474
25.6562 | -35.8215 | -39.4799
25.6623 | -35.822 | -39.4856
25.6684 | -35.8225 | -39.4915
25.6745 | -35.823 | -39.4974
25.6806 | -35.8234 | -39.5029
25.6867 | -35.8239 | -39.5088
25.6928 | -35.8244 | -39.5147
25.6989 | -35.8248 | -39.5207
25.705 | -35.8253 | -39.5267
25.7111 | -35.8258 | -39.5326
25.7173 | -35.8262 | -39.5383
25.7234 | -35.8267 | -39.5442
25.7295 | -35.8272 | -39.5501
25.7356 | -35.8276 | -39.5561
25.7417 | -35.8281 | -39.5617
25.7478 | -35.8286 | -39.5675
25.7539 | -35.829 | -39.5732
25.76 | -35.8295 | -39.5792
25.7661 | -35.83 | -39.5849
25.7722 | -35.8305 | -39.5906
25.7783 | -35.831 | -39.5965
25.7844 | -35.8314 | -39.6021
25.7905 | -35.8319 | -39.6079
25.7966 | -35.8323 | -39.6136
25.8027 | -35.8327 | -39.6194
25.8088 | -35.8332 | -39.6253
25.8149 | -35.8337 | -39.631
25.821 | -35.8341 | -39.6367
25.8271 | -35.8346 | -39.6427
25.8332 | -35.8351 | -39.6485
25.8393 | -35.8355 | -39.6544
25.8454 | -35.836 | -39.6601
25.8515 | -35.8365 | -39.6659
25.8576 | -35.837 | -39.6715
25.8638 | -35.8375 | -39.6774
25.8699 | -35.8379 | -39.6832
25.876 | -35.8384 | -39.6891
25.8821 | -35.8389 | -39.6949
25.8882 | -35.8393 | -39.7006
25.8943 | -35.8398 | -39.7065
25.9004 | -35.8403 | -39.7124
25.9065 | -35.8408 | -39.718
25.9126 | -35.8412 | -39.7238
25.9187 | -35.8417 | -39.7295
25.9248 | -35.8422 | -39.7353
25.9309 | -35.8426 | -39.7411
25.937 | -35.8431 | -39.7471
25.9431 | -35.8436 | -39.7529
25.9492 | -35.844 | -39.7586
25.9553 | -35.8445 | -39.7644
25.9614 | -35.845 | -39.7702
25.9675 | -35.8454 | -39.7756
25.9736 | -35.8459 | -39.7815
25.9797 | -35.8464 | -39.7872
25.9858 | -35.8468 | -39.7929
25.9919 | -35.8473 | -39.7989
25.998 | -35.8477 | -39.8044
26.0042 | -35.8482 | -39.8101
26.0103 | -35.8486 | -39.8157
26.0164 | -35.849 | -39.8215
26.0225 | -35.8495 | -39.8273
26.0286 | -35.85 | -39.8334
26.0347 | -35.8504 | -39.8388
26.0408 | -35.8509 | -39.8446
26.0469 | -35.8513 | -39.8503
26.053 | -35.8518 | -39.8561
26.0591 | -35.8523 | -39.862
26.0652 | -35.8527 | -39.8677
26.0713 | -35.8532 | -39.8735
26.0774 | -35.8537 | -39.8792
26.0835 | -35.8541 | -39.8851
26.0896 | -35.8546 | -39.8907
26.0957 | -35.855 | -39.8964
26.1018 | -35.8555 | -39.902
26.1079 | -35.856 | -39.9074
26.114 | -35.8564 | -39.9132
26.1201 | -35.8569 | -39.9191
26.1262 | -35.8573 | -39.9252
26.1323 | -35.8578 | -39.9307
26.1384 | -35.8582 | -39.9364
26.1445 | -35.8587 | -39.9422
26.1507 | -35.8592 | -39.9479
26.1568 | -35.8596 | -39.9536
26.1629 | -35.8601 | -39.9592
26.169 | -35.8605 | -39.9649
26.1751 | -35.8609 | -39.9708
26.1812 | -35.8614 | -39.9764
26.1873 | -35.8619 | -39.9822
26.1934 | -35.8624 | -39.9878
26.1995 | -35.8629 | -39.9937
26.2056 | -35.8633 | -39.9993
26.2117 | -35.8638 | -40.0052
26.2178 | -35.8642 | -40.0109
26.2239 | -35.8647 | -40.0166
26.23 | -35.8652 | -40.0225
26.2361 | -35.8657 | -40.0282
26.2422 | -35.8661 | -40.0338
26.2483 | -35.8666 | -40.0392
26.2544 | -35.867 | -40.0448
26.2605 | -35.8675 | -40.0506
26.2666 | -35.868 | -40.0564
26.2727 | -35.8684 | -40.0623
26.2788 | -35.8688 | -40.0677
26.2849 | -35.8693 | -40.0735
26.2911 | -35.8697 | -40.0792
26.2972 | -35.8702 | -40.085
26.3033 | -35.8707 | -40.0906
26.3094 | -35.8711 | -40.0962
26.3155 | -35.8715 | -40.1015
26.3216 | -35.872 | -40.1073
26.3277 | -35.8724 | -40.1126
26.3338 | -35.8728 | -40.1182
26.3399 | -35.8733 | -40.1237
26.346 | -35.8737 | -40.1291
26.3521 | -35.8742 | -40.135
26.3582 | -35.8747 | -40.1407
26.3643 | -35.8751 | -40.1464
26.3704 | -35.8756 | -40.1523
26.3765 | -35.8761 | -40.1581
26.3826 | -35.8766 | -40.1636
26.3887 | -35.877 | -40.1693
26.3948 | -35.8774 | -40.1749
26.4009 | -35.8779 | -40.1806
26.407 | -35.8783 | -40.1863
26.4131 | -35.8788 | -40.1919
26.4192 | -35.8792 | -40.1977
26.4253 | -35.8796 | -40.2032
26.4314 | -35.8801 | -40.2088
26.4376 | -35.8806 | -40.2147
26.4437 | -35.881 | -40.2204
26.4498 | -35.8814 | -40.2263
26.4559 | -35.8819 | -40.2319
26.462 | -35.8823 | -40.2375
26.4681 | -35.8828 | -40.2431
26.4742 | -35.8832 | -40.249
26.4803 | -35.8837 | -40.2549
26.4864 | -35.8842 | -40.2606
26.4925 | -35.8846 | -40.2664
26.4986 | -35.885 | -40.2723
26.5047 | -35.8855 | -40.2782
26.5108 | -35.886 | -40.2837
26.5169 | -35.8864 | -40.2892
26.523 | -35.8868 | -40.2948
26.5291 | -35.8873 | -40.3002
26.5352 | -35.8877 | -40.3059
26.5413 | -35.8882 | -40.3116
26.5474 | -35.8886 | -40.3173
26.5535 | -35.8891 | -40.3231
26.5596 | -35.8895 | -40.329
26.5657 | -35.8899 | -40.3347
26.5718 | -35.8904 | -40.3405
26.578 | -35.8908 | -40.3463
26.5841 | -35.8913 | -40.352
26.5902 | -35.8917 | -40.3575
26.5963 | -35.8922 | -40.3633
26.6024 | -35.8926 | -40.3691
26.6085 | -35.893 | -40.3746
26.6146 | -35.8935 | -40.3801
26.6207 | -35.8939 | -40.386
26.6268 | -35.8944 | -40.392
26.6329 | -35.8948 | -40.3977
26.639 | -35.8952 | -40.4034
26.6451 | -35.8956 | -40.409
26.6512 | -35.896 | -40.4146
26.6573 | -35.8965 | -40.4203
26.6634 | -35.8969 | -40.4257
26.6695 | -35.8973 | -40.4317
26.6756 | -35.8977 | -40.4377
26.6817 | -35.8982 | -40.4434
26.6878 | -35.8986 | -40.4492
26.6939 | -35.899 | -40.4549
26.7 | -35.8995 | -40.4604
26.7061 | -35.8999 | -40.466
26.7122 | -35.9004 | -40.4716
26.7183 | -35.9008 | -40.4772
26.7245 | -35.9013 | -40.483
26.7306 | -35.9018 | -40.4886
26.7367 | -35.9022 | -40.4943
26.7428 | -35.9026 | -40.4998
26.7489 | -35.9031 | -40.5055
26.755 | -35.9036 | -40.5109
26.7611 | -35.904 | -40.5166
26.7672 | -35.9044 | -40.5223
26.7733 | -35.9049 | -40.5276
26.7794 | -35.9053 | -40.5333
26.7855 | -35.9058 | -40.539
26.7916 | -35.9062 | -40.5446
26.7977 | -35.9067 | -40.5503
26.8038 | -35.9071 | -40.5562
26.8099 | -35.9076 | -40.5619
26.816 | -35.908 | -40.5674
26.8221 | -35.9084 | -40.5732
26.8282 | -35.9088 | -40.5786
26.8343 | -35.9093 | -40.5845
26.8404 | -35.9097 | -40.59
26.8465 | -35.9102 | -40.5958
26.8526 | -35.9106 | -40.6014
26.8587 | -35.9111 | -40.6071
26.8649 | -35.9115 | -40.6126
26.871 | -35.9119 | -40.6182
26.8771 | -35.9124 | -40.6237
26.8832 | -35.9128 | -40.6292
26.8893 | -35.9132 | -40.6351
26.8954 | -35.9137 | -40.6409
26.9015 | -35.9141 | -40.6465
26.9076 | -35.9145 | -40.6524
26.9137 | -35.9149 | -40.6582
26.9198 | -35.9154 | -40.6641
26.9259 | -35.9158 | -40.6697
26.932 | -35.9162 | -40.6751
26.9381 | -35.9167 | -40.6809
26.9442 | -35.9171 | -40.6863
26.9503 | -35.9175 | -40.692
26.9564 | -35.918 | -40.6979
26.9625 | -35.9184 | -40.7034
26.9686 | -35.9189 | -40.7093
26.9747 | -35.9193 | -40.715
26.9808 | -35.9198 | -40.7204
26.9869 | -35.9202 | -40.7259
26.993 | -35.9206 | -40.7313
26.9991 | -35.9211 | -40.7369
27.0052 | -35.9215 | -40.7426
27.0114 | -35.9219 | -40.7484
27.0175 | -35.9224 | -40.7536
27.0236 | -35.9228 | -40.7593
27.0297 | -35.9232 | -40.765
27.0358 | -35.9237 | -40.7708
27.0419 | -35.9241 | -40.7765
27.048 | -35.9246 | -40.7822
27.0541 | -35.925 | -40.7879
27.0602 | -35.9254 | -40.7935
27.0663 | -35.9259 | -40.7994
27.0724 | -35.9263 | -40.8048
27.0785 | -35.9268 | -40.8106
27.0846 | -35.9272 | -40.8159
27.0907 | -35.9276 | -40.8215
27.0968 | -35.9281 | -40.8272
27.1029 | -35.9285 | -40.8326
27.109 | -35.929 | -40.8382
27.1151 | -35.9294 | -40.8437
27.1212 | -35.9298 | -40.849
27.1273 | -35.9303 | -40.8545
27.1334 | -35.9307 | -40.8603
27.1395 | -35.9312 | -40.8661
27.1456 | -35.9316 | -40.8718
27.1518 | -35.932 | -40.8773
27.1579 | -35.9324 | -40.8829
27.164 | -35.9328 | -40.8885
27.1701 | -35.9333 | -40.8941
27.1762 | -35.9337 | -40.8998
27.1823 | -35.9341 | -40.9056
27.1884 | -35.9345 | -40.9111
27.1945 | -35.9349 | -40.9166
27.2006 | -35.9354 | -40.922
27.2067 | -35.9358 | -40.9276
27.2128 | -35.9362 | -40.9333
27.2189 | -35.9366 | -40.9389
27.225 | -35.937 | -40.9446
27.2311 | -35.9375 | -40.9501
27.2372 | -35.938 | -40.9558
27.2433 | -35.9384 | -40.9615
27.2494 | -35.9388 | -40.967
27.2555 | -35.9392 | -40.9726
27.2616 | -35.9396 | -40.978
27.2677 | -35.94 | -40.9838
27.2738 | -35.9405 | -40.9892
27.2799 | -35.9409 | -40.9949
27.286 | -35.9413 | -41.0002
27.2921 | -35.9418 | -41.0056
27.2983 | -35.9422 | -41.0109
27.3044 | -35.9426 | -41.0165
27.3105 | -35.943 | -41.0223
27.3166 | -35.9434 | -41.0278
27.3227 | -35.9439 | -41.0332
27.3288 | -35.9443 | -41.0386
27.3349 | -35.9446 | -41.0438
27.341 | -35.9451 | -41.0496
27.3471 | -35.9455 | -41.0554
27.3532 | -35.9459 | -41.0609
27.3593 | -35.9464 | -41.0666
27.3654 | -35.9469 | -41.0721
27.3715 | -35.9473 | -41.0778
27.3776 | -35.9477 | -41.0834
27.3837 | -35.9481 | -41.0891
27.3898 | -35.9486 | -41.0946
27.3959 | -35.949 | -41.1001
27.402 | -35.9494 | -41.1053
27.4081 | -35.9499 | -41.1109
27.4142 | -35.9503 | -41.1164
27.4203 | -35.9507 | -41.122
27.4264 | -35.9511 | -41.1276
27.4325 | -35.9515 | -41.1327
27.4387 | -35.952 | -41.1381
27.4448 | -35.9524 | -41.1439
27.4509 | -35.9528 | -41.1496
27.457 | -35.9532 | -41.1548
27.4631 | -35.9536 | -41.1603
27.4692 | -35.954 | -41.1654
27.4753 | -35.9544 | -41.1705
27.4814 | -35.9548 | -41.1759
27.4875 | -35.9552 | -41.1816
27.4936 | -35.9557 | -41.1873
27.4997 | -35.9561 | -41.1927
27.5058 | -35.9565 | -41.1984
27.5119 | -35.9569 | -41.204
27.518 | -35.9574 | -41.2094
27.5241 | -35.9578 | -41.2148
27.5302 | -35.9583 | -41.2203
27.5363 | -35.9587 | -41.226
27.5424 | -35.9591 | -41.2312
27.5485 | -35.9595 | -41.2366
27.5546 | -35.9599 | -41.2419
27.5607 | -35.9604 | -41.2472
27.5668 | -35.9608 | -41.2527
27.5729 | -35.9612 | -41.2581
27.5791 | -35.9616 | -41.2634
27.5852 | -35.9621 | -41.2688
27.5913 | -35.9625 | -41.2745
27.5974 | -35.9629 | -41.2802
27.6035 | -35.9633 | -41.2856
27.6096 | -35.9637 | -41.291
27.6157 | -35.9642 | -41.2963
27.6218 | -35.9646 | -41.302
27.6279 | -35.965 | -41.3076
27.634 | -35.9654 | -41.313
27.6401 | -35.9658 | -41.3185
27.6462 | -35.9663 | -41.3238
27.6523 | -35.9666 | -41.3286
27.6584 | -35.9671 | -41.3341
27.6645 | -35.9675 | -41.3396
27.6706 | -35.9679 | -41.3453
27.6767 | -35.9684 | -41.3509
27.6828 | -35.9688 | -41.3563
27.6889 | -35.9692 | -41.3616
27.695 | -35.9696 | -41.3669
27.7011 | -35.97 | -41.3724
27.7072 | -35.9704 | -41.3779
27.7133 | -35.9709 | -41.3835
27.7194 | -35.9713 | -41.3885
27.7256 | -35.9717 | -41.3939
27.7317 | -35.9721 | -41.3993
27.7378 | -35.9725 | -41.4048
27.7439 | -35.9729 | -41.4103
27.75 | -35.9733 | -41.4159
27.7561 | -35.9738 | -41.4211
27.7622 | -35.9742 | -41.4266
27.7683 | -35.9746 | -41.4321
27.7744 | -35.975 | -41.4376
27.7805 | -35.9754 | -41.4429
27.7866 | -35.9758 | -41.4482
27.7927 | -35.9762 | -41.4537
27.7988 | -35.9766 | -41.4591
27.8049 | -35.9771 | -41.4645
27.811 | -35.9775 | -41.4701
27.8171 | -35.9779 | -41.4757
27.8232 | -35.9783 | -41.481
27.8293 | -35.9787 | -41.4865
27.8354 | -35.9792 | -41.4921
27.8415 | -35.9796 | -41.4977
27.8476 | -35.98 | -41.5031
27.8537 | -35.9804 | -41.5086
27.8598 | -35.9808 | -41.5139
27.866 | -35.9812 | -41.5191
27.8721 | -35.9817 | -41.5248
27.8782 | -35.9821 | -41.5303
27.8843 | -35.9825 | -41.5359
27.8904 | -35.9829 | -41.5414
27.8965 | -35.9834 | -41.5468
27.9026 | -35.9838 | -41.5522
27.9087 | -35.9842 | -41.5576
27.9148 | -35.9846 | -41.5629
27.9209 | -35.985 | -41.5685
27.927 | -35.9854 | -41.5738
27.9331 | -35.9859 | -41.5792
27.9392 | -35.9863 | -41.5846
27.9453 | -35.9867 | -41.59
27.9514 | -35.9872 | -41.5954
27.9575 | -35.9876 | -41.6009
27.9636 | -35.9879 | -41.6064
27.9697 | -35.9883 | -41.6115
27.9758 | -35.9887 | -41.6171
27.9819 | -35.9891 | -41.6228
27.988 | -35.9895 | -41.6282
27.9941 | -35.99 | -41.6339
28.0002 | -35.9904 | -41.6393
28.0063 | -35.9908 | -41.6448
28.0125 | -35.9912 | -41.6505
28.0186 | -35.9915 | -41.6562
28.0247 | -35.992 | -41.6616
28.0308 | -35.9924 | -41.6669
28.0369 | -35.9928 | -41.6718
28.043 | -35.9932 | -41.6777
28.0491 | -35.9937 | -41.6832
28.0552 | -35.9941 | -41.6885
28.0613 | -35.9945 | -41.6944
28.0674 | -35.9949 | -41.6994
28.0735 | -35.9953 | -41.7045

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.6936 | -47.1059
21.4908 | -35.802 | -38.8457
16.4948 | -35.0254 | -31.174
12.6602 | -34.57 | -24.391
9.7171 | -34.2902 | -18.9612
7.4582 | -34.0914 | -14.5127
5.7244 | -33.9487 | -10.8759
4.3936 | -33.8609 | -7.9554
3.3722 | -33.8346 | -5.6346
2.5883 | -33.8278 | -3.8733
1.9866 | -33.8127 | -2.462
1.5248 | -33.8191 | -1.337
1.1703 | -33.8184 | -0.58109
1 | -33.8137 | -0.2114
0.89822 | -33.8055 | 0.092981
0.68941 | -33.7924 | 0.73815
0.52919 | -33.7962 | 1.242
0.40612 | -33.793 | 1.8974
0.31171 | -33.7903 | 2.3832
0.23928 | -33.8102 | 3.126
0.18358 | -33.8033 | 3.5739
0.14091 | -33.8254 | 4.457
0.10819 | -33.8238 | 5.6303
0.083017 | -33.8493 | 7.1109
0.063705 | -33.9021 | 8.985
0.048923 | -33.9721 | 11.4922
0.037479 | -34.0723 | 14.6888
0.028801 | -34.2508 | 18.7325
0.022078 | -34.5635 | 24.1427
0.017023 | -35.0487 | 30.573
0.013018 | -35.852 | 38.2446
0.010014 | -36.9647 | 46.1829

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003704 | 7.3317
0.007407 | -26.5802
0.011111 | -33.7097
0.014815 | -38.27
0.018519 | -40.7533
0.022222 | -44.2055
0.025926 | -45.6987
0.02963 | -46.5809
0.033333 | -47.8111
0.037037 | -48.4857
0.040741 | -49.16
0.044444 | -51.2917
0.048148 | -52.9218
0.051852 | -55.1137
0.055556 | -55.6757
0.059259 | -55.4602
0.062963 | -55.5389
0.066667 | -56.5635
0.07037 | -57.6592
0.074074 | -57.8041
0.077778 | -58.7027
0.081481 | -59.767
0.085185 | -60.5556
0.088889 | -61.6126
0.092593 | -61.4192
0.096296 | -63.0039
0.1 | -63.5296
0.1037 | -62.9651
0.10741 | -63.6165
0.11111 | -63.7196
0.11482 | -64.5622
0.11852 | -64.8432
0.12222 | -64.8628
0.12593 | -64.9602
0.12963 | -66.4743
0.13333 | -67.0951
0.13704 | -66.8871
0.14074 | -66.1488
0.14444 | -66.0906
0.14815 | -66.5692
0.15185 | -68.1495
0.15556 | -67.8121
0.15926 | -67.5428
0.16296 | -68.2554
0.16667 | -68.1272
0.17037 | -68.4585
0.17407 | -69.1292
0.17778 | -68.0256
0.18148 | -68.4895
0.18518 | -68.6542
0.18889 | -68.9661
0.19259 | -70.7736
0.1963 | -70.7498
0.2 | -71.049
0.2037 | -70.5787
0.20741 | -70.128
0.21111 | -70.7381
0.21482 | -70.6563
0.21852 | -70.7008
0.22222 | -70.522
0.22593 | -71.4943
0.22963 | -72.2717
0.23333 | -72.7535
0.23704 | -73.2716
0.24074 | -72.8823
0.24444 | -73.1288
0.24815 | -72.7952
0.25185 | -72.2776
0.25556 | -72.8369
0.25926 | -72.5828
0.26296 | -73.1076
0.26667 | -73.2967
0.27037 | -73.3735
0.27407 | -73.5187
0.27778 | -73.7786
0.28148 | -73.336
0.28519 | -72.4846
0.28889 | -72.8855
0.29259 | -74.5383
0.2963 | -74.9689
0.3 | -75.0268
0.3037 | -74.8697
0.30741 | -74.8211
0.31111 | -75.4046
0.31482 | -76.3355
0.31852 | -75.8069
0.32222 | -75.3064
0.32593 | -75.5315
0.32963 | -76.1503
0.33333 | -75.7624
0.33704 | -75.6671
0.34074 | -76.7023
0.34444 | -76.6838
0.34815 | -76.1702
0.35185 | -74.8784
0.35556 | -74.8795
0.35926 | -75.2905
0.36296 | -75.9378
0.36667 | -76.8613
0.37037 | -76.6926
0.37407 | -76.3456
0.37778 | -76.8533
0.38148 | -77.2593
0.38518 | -77.5388
0.38889 | -77.3666
0.39259 | -76.8268
0.3963 | -75.9977
0.4 | -75.9832
0.4037 | -77.0821
0.40741 | -77.189
0.41111 | -77.1462
0.41481 | -77.8918
0.41852 | -78.7167
0.42222 | -78.3822
0.42593 | -78.0613
0.42963 | -77.8142
0.43333 | -78.4871
0.43704 | -79.538
0.44074 | -78.7716
0.44444 | -77.5296
0.44815 | -78.164
0.45185 | -78.106
0.45556 | -77.6951
0.45926 | -79.9739
0.46296 | -79.7334
0.46667 | -79.3969
0.47037 | -78.9006
0.47407 | -79.2706
0.47778 | -79.4061
0.48148 | -79.6838
0.48518 | -80.151
0.48889 | -80.1442
0.49259 | -79.3056
0.4963 | -78.8849
0.5 | -79.1732
0.5037 | -78.9314
0.50741 | -78.8864
0.51111 | -79.5334
0.51482 | -80.1048
0.51852 | -80.0218
0.52222 | -80.1183
0.52593 | -79.8089
0.52963 | -79.9784
0.53333 | -79.9475
0.53704 | -79.776
0.54074 | -79.9367
0.54444 | -80.1275
0.54815 | -79.9963
0.55185 | -80.0863
0.55556 | -80.3636
0.55926 | -80.2721
0.56296 | -80.3349
0.56667 | -80.7719
0.57037 | -81.1477
0.57407 | -81.4318
0.57778 | -81.6735
0.58148 | -81.7315
0.58518 | -81.6428
0.58889 | -81.7032
0.59259 | -82.1474
0.5963 | -82.084
0.6 | -81.6736
0.6037 | -81.5251
0.60741 | -81.8341
0.61111 | -81.9358
0.61482 | -81.696
0.61852 | -81.7085
0.62222 | -81.5669
0.62593 | -81.5564
0.62963 | -81.7877
0.63333 | -81.999
0.63704 | -82.1298
0.64074 | -82.2525
0.64444 | -82.2251
0.64815 | -81.8457
0.65185 | -81.494
0.65556 | -81.6522
0.65926 | -81.6511
0.66296 | -81.9602
0.66667 | -82.8049
0.67037 | -83.5466
0.67407 | -83.3519
0.67778 | -82.6019
0.68148 | -82.1747
0.68519 | -82.3268
0.68889 | -82.703
0.69259 | -82.9541
0.6963 | -83.023
0.7 | -82.8081
0.7037 | -82.6805
0.70741 | -82.9586
0.71111 | -83.0456
0.71481 | -82.6951
0.71852 | -82.7803
0.72222 | -83.2399
0.72593 | -83.1511
0.72963 | -82.8193
0.73333 | -82.7456
0.73704 | -82.9714
0.74074 | -83.4604
0.74444 | -84.073
0.74815 | -84.3197
0.75185 | -84.5401
0.75556 | -84.4584
0.75926 | -84.0226
0.76296 | -83.7528
0.76667 | -83.6244
0.77037 | -83.5766
0.77407 | -83.9645
0.77778 | -84.3551
0.78148 | -84.4489
0.78519 | -84.5476
0.78889 | -84.452
0.79259 | -84.0577
0.7963 | -83.8981
0.8 | -83.8958
0.8037 | -84.0781
0.80741 | -84.6426
0.81111 | -84.903
0.81481 | -84.5536
0.81852 | -83.7887
0.82222 | -83.3922
0.82593 | -83.7834
0.82963 | -84.4313
0.83333 | -85.1297
0.83704 | -85.9362
0.84074 | -85.8776
0.84444 | -85.1301
0.84815 | -84.6396
0.85185 | -84.3803
0.85556 | -84.0834
0.85926 | -84.0212
0.86296 | -84.3446
0.86667 | -84.6458
0.87037 | -84.4087
0.87407 | -84.2846
0.87778 | -84.5712
0.88148 | -84.4173
0.88518 | -84.28
0.88889 | -84.671
0.89259 | -84.9907
0.8963 | -85.0798
0.9 | -85.071
0.9037 | -85.013
0.90741 | -84.653
0.91111 | -84.0798
0.91482 | -83.8899
0.91852 | -84.2838
0.92222 | -84.8232
0.92593 | -85.4234
0.92963 | -85.9201
0.93333 | -86.0899
0.93704 | -86.2123
0.94074 | -86.3034
0.94444 | -86.2858
0.94815 | -86.2775
0.95185 | -86.2991
0.95556 | -86.253
0.95926 | -86.3085
0.96296 | -86.1986
0.96667 | -85.7042
0.97037 | -85.042
0.97407 | -85.1397
0.97778 | -86.1913
0.98148 | -87.1181
0.98518 | -86.9986
0.98889 | -85.908
0.99259 | -82.2692
0.9963 | -77.7969
1 | -75.9862
1.0037 | -78.0576
1.0074 | -82.9554
1.0111 | -86.5316
1.0148 | -86.8814
1.0185 | -86.8533
1.0222 | -86.6468
1.0259 | -86.5221
1.0296 | -87.1983
1.0333 | -88.0175
1.037 | -87.9735
1.0407 | -87.3001
1.0444 | -86.9046
1.0481 | -87.153
1.0519 | -87.1261
1.0556 | -86.915
1.0593 | -87.1325
1.063 | -87.4094
1.0667 | -87.3069
1.0704 | -87.2958
1.0741 | -87.2758
1.0778 | -87.2247
1.0815 | -87.219
1.0852 | -87.3341
1.0889 | -87.5317
1.0926 | -87.6354
1.0963 | -87.5031
1.1 | -87.1917
1.1037 | -86.8971
1.1074 | -86.7764
1.1111 | -86.7892
1.1148 | -86.7757
1.1185 | -86.6714
1.1222 | -86.6011
1.1259 | -86.6392
1.1296 | -86.6352
1.1333 | -86.5725
1.137 | -86.6622
1.1407 | -86.9615
1.1444 | -87.3581
1.1481 | -87.7266
1.1519 | -88.095
1.1556 | -88.41
1.1593 | -88.2574
1.163 | -87.6776
1.1667 | -87.2478
1.1704 | -87.1982
1.1741 | -87.431
1.1778 | -87.8136
1.1815 | -88.2502
1.1852 | -88.479
1.1889 | -88.2339
1.1926 | -87.8516
1.1963 | -87.8636
1.2 | -88.2915
1.2037 | -88.7171
1.2074 | -88.7998
1.2111 | -88.6927
1.2148 | -88.6027
1.2185 | -88.4973
1.2222 | -88.3535
1.2259 | -88.2123
1.2296 | -88.0209
1.2333 | -87.8506
1.237 | -87.8561
1.2407 | -88.07
1.2444 | -88.4777
1.2481 | -88.9134
1.2519 | -89.0473
1.2556 | -88.8369
1.2593 | -88.5852
1.263 | -88.5317
1.2667 | -88.7303
1.2704 | -89.071
1.2741 | -89.2964
1.2778 | -89.2877
1.2815 | -89.1769
1.2852 | -88.9345
1.2889 | -88.529
1.2926 | -88.2112
1.2963 | -88.1716
1.3 | -88.348
1.3037 | -88.6392
1.3074 | -89.0797
1.3111 | -89.6787
1.3148 | -90.1411
1.3185 | -89.9756
1.3222 | -89.2918
1.3259 | -88.7036
1.3296 | -88.5355
1.3333 | -88.6637
1.337 | -88.7601
1.3407 | -88.6535
1.3444 | -88.4302
1.3481 | -88.2441
1.3519 | -88.227
1.3556 | -88.4548
1.3593 | -88.8441
1.363 | -89.1854
1.3667 | -89.3893
1.3704 | -89.5645
1.3741 | -89.6985
1.3778 | -89.7159
1.3815 | -89.6983
1.3852 | -89.7742
1.3889 | -89.9487
1.3926 | -90.131
1.3963 | -90.2346
1.4 | -90.2017
1.4037 | -90.1111
1.4074 | -90.1118
1.4111 | -90.1069
1.4148 | -89.8608
1.4185 | -89.3845
1.4222 | -89.0045
1.4259 | -88.9557
1.4296 | -89.1749
1.4333 | -89.4846
1.437 | -89.7946
1.4407 | -90.1184
1.4444 | -90.4232
1.4481 | -90.5456
1.4519 | -90.3713
1.4556 | -89.9929
1.4593 | -89.6731
1.463 | -89.6217
1.4667 | -89.7747
1.4704 | -89.8952
1.4741 | -89.8639
1.4778 | -89.8033
1.4815 | -89.8108
1.4852 | -89.7721
1.4889 | -89.6014
1.4926 | -89.4015
1.4963 | -89.3092
1.5 | -89.3706
1.5037 | -89.5019
1.5074 | -89.5786
1.5111 | -89.6318
1.5148 | -89.7477
1.5185 | -89.8727
1.5222 | -89.9211
1.5259 | -89.9364
1.5296 | -90.0042
1.5333 | -90.1695
1.537 | -90.4204
1.5407 | -90.6767
1.5444 | -90.8943
1.5481 | -91.147
1.5519 | -91.3637
1.5556 | -91.3168
1.5593 | -91.0081
1.563 | -90.6946
1.5667 | -90.6132
1.5704 | -90.7751
1.5741 | -90.9919
1.5778 | -91.099
1.5815 | -91.0138
1.5852 | -90.7748
1.5889 | -90.6097
1.5926 | -90.6935
1.5963 | -90.9106
1.6 | -90.8567
1.6037 | -90.7282
1.6074 | -90.4988
1.6111 | -90.2962
1.6148 | -90.1915
1.6185 | -90.1977
1.6222 | -90.3031
1.6259 | -90.5108
1.6296 | -90.7884
1.6333 | -91.0571
1.637 | -91.2476
1.6407 | -91.2909
1.6444 | -91.1641
1.6481 | -90.9269
1.6519 | -90.6913
1.6556 | -90.5871
1.6593 | -90.6734
1.663 | -90.8971
1.6667 | -91.1298
1.6704 | -91.2655
1.6741 | -91.303
1.6778 | -91.3045
1.6815 | -91.2691
1.6852 | -91.1559
1.6889 | -90.9531
1.6926 | -90.7194
1.6963 | -90.5409
1.7 | -90.5066
1.7037 | -90.6215
1.7074 | -90.7711
1.7111 | -90.8552
1.7148 | -90.8846
1.7185 | -90.9547
1.7222 | -91.1003
1.7259 | -91.2508
1.7296 | -91.3448
1.7333 | -91.3875
1.737 | -91.454
1.7407 | -91.5574
1.7444 | -91.5984
1.7481 | -91.4902
1.7519 | -91.2678
1.7556 | -91.0589
1.7593 | -90.9862
1.763 | -91.117
1.7667 | -91.4473
1.7704 | -91.8203
1.7741 | -92.0653
1.7778 | -92.1623
1.7815 | -92.1988
1.7852 | -92.2316
1.7889 | -92.2903
1.7926 | -92.3491
1.7963 | -92.4101
1.8 | -92.4726
1.8037 | -92.514
1.8074 | -92.4929
1.8111 | -92.3936
1.8148 | -92.2111
1.8185 | -91.8845
1.8222 | -91.4181
1.8259 | -90.9753
1.8296 | -90.7648
1.8333 | -90.8761
1.837 | -91.2465
1.8407 | -91.7213
1.8444 | -92.1091
1.8481 | -92.213
1.8519 | -92.0197
1.8556 | -91.7474
1.8593 | -91.6305
1.863 | -91.7304
1.8667 | -91.9999
1.8704 | -92.3594
1.8741 | -92.7288
1.8778 | -92.9587
1.8815 | -92.9262
1.8852 | -92.6589
1.8889 | -92.3249
1.8926 | -92.1284
1.8963 | -92.1626
1.9 | -92.2984
1.9037 | -92.3407
1.9074 | -92.2608
1.9111 | -92.2071
1.9148 | -92.2658
1.9185 | -92.4169
1.9222 | -92.5331
1.9259 | -92.4897
1.9296 | -92.3076
1.9333 | -92.1633
1.937 | -92.1461
1.9407 | -92.2459
1.9444 | -92.3955
1.9481 | -92.5162
1.9519 | -92.5569
1.9556 | -92.5209
1.9593 | -92.4574
1.963 | -92.4206
1.9667 | -92.3873
1.9704 | -92.3262
1.9741 | -92.2279
1.9778 | -92.1296
1.9815 | -91.8142
1.9852 | -90.3
1.9889 | -87.9705
1.9926 | -85.5658
1.9963 | -83.8776
2 | -83.2734
2.0037 | -83.8402
2.0074 | -85.4962
2.0111 | -87.8998
2.0148 | -90.2972
2.0185 | -91.9616
2.0222 | -92.4131
2.0259 | -92.592
2.0296 | -92.7117
2.0333 | -92.6562
2.037 | -92.471
2.0407 | -92.307
2.0444 | -92.3313
2.0481 | -92.5908
2.0519 | -93.0135
2.0556 | -93.4301
2.0593 | -93.6538
2.063 | -93.6225
2.0667 | -93.4588
2.0704 | -93.2999
2.0741 | -93.1777
2.0778 | -93.0602
2.0815 | -92.9387
2.0852 | -92.8258
2.0889 | -92.7557
2.0926 | -92.7775
2.0963 | -92.9118
2.1 | -93.0772
2.1037 | -93.1561
2.1074 | -93.1059
2.1111 | -92.9976
2.1148 | -92.924
2.1185 | -92.9215
2.1222 | -92.9688
2.1259 | -93.0139
2.1296 | -93.0336
2.1333 | -93.0165
2.137 | -92.951
2.1407 | -92.9117
2.1444 | -92.8837
2.1481 | -92.8752
2.1519 | -92.8746
2.1556 | -92.8692
2.1593 | -92.8805
2.163 | -92.9471
2.1667 | -93.0913
2.1704 | -93.3169
2.1741 | -93.5892
2.1778 | -93.8448
2.1815 | -94.0003
2.1852 | -94.0037
2.1889 | -93.8552
2.1926 | -93.6386
2.1963 | -93.4689
2.2 | -93.414
2.2037 | -93.5028
2.2074 | -93.6936
2.2111 | -93.9271
2.2148 | -94.1452
2.2185 | -94.2893
2.2222 | -94.3221
2.2259 | -94.2235
2.2296 | -94.0067
2.2333 | -93.7407
2.237 | -93.5103
2.2407 | -93.351
2.2444 | -93.2553
2.2481 | -93.204
2.2519 | -93.19
2.2556 | -93.2088
2.2593 | -93.2812
2.263 | -93.3995
2.2667 | -93.5288
2.2704 | -93.6459
2.2741 | -93.7268
2.2778 | -93.7663
2.2815 | -93.7942
2.2852 | -93.8234
2.2889 | -93.8517
2.2926 | -93.8555
2.2963 | -93.8426
2.3 | -93.7966
2.3037 | -93.7451
2.3074 | -93.7364
2.3111 | -93.7925
2.3148 | -93.903
2.3185 | -94.0243
2.3222 | -94.1166
2.3259 | -94.1807
2.3296 | -94.2455
2.3333 | -94.3236
2.337 | -94.3808
2.3407 | -94.3912
2.3444 | -94.3299
2.3481 | -94.2171
2.3519 | -94.1183
2.3556 | -94.0898
2.3593 | -94.1447
2.363 | -94.2354
2.3667 | -94.3163
2.3704 | -94.3599
2.3741 | -94.363
2.3778 | -94.3535
2.3815 | -94.3446
2.3852 | -94.351
2.3889 | -94.358
2.3926 | -94.3225
2.3963 | -94.2161
2.4 | -94.0511
2.4037 | -93.8869
2.4074 | -93.7443
2.4111 | -93.6476
2.4148 | -93.6158
2.4185 | -93.6642
2.4222 | -93.7989
2.4259 | -94.0154
2.4296 | -94.3075
2.4333 | -94.6332
2.437 | -94.9042
2.4407 | -95.0096
2.4444 | -94.8907
2.4481 | -94.593
2.4519 | -94.2195
2.4556 | -93.9029
2.4593 | -93.7338
2.463 | -93.7543
2.4667 | -93.9348
2.4704 | -94.221
2.4741 | -94.5301
2.4778 | -94.7678
2.4815 | -94.8597
2.4852 | -94.7893
2.4889 | -94.6047
2.4926 | -94.3912
2.4963 | -94.2313
2.5 | -94.1879
2.5037 | -94.261
2.5074 | -94.4033
2.5111 | -94.5555
2.5148 | -94.682
2.5185 | -94.7566
2.5222 | -94.7778
2.5259 | -94.7584
2.5296 | -94.7266
2.5333 | -94.7109
2.537 | -94.7486
2.5407 | -94.8503
2.5444 | -95.0063
2.5481 | -95.1683
2.5519 | -95.2617
2.5556 | -95.2467
2.5593 | -95.158
2.563 | -95.0363
2.5667 | -94.9116
2.5704 | -94.7865
2.5741 | -94.6506
2.5778 | -94.4613
2.5815 | -94.2292
2.5852 | -94.0283
2.5889 | -93.9402
2.5926 | -94.0011
2.5963 | -94.1752
2.6 | -94.3933
2.6037 | -94.6082
2.6074 | -94.8016
2.6111 | -94.9542
2.6148 | -95.0371
2.6185 | -95.0397
2.6222 | -94.9628
2.6259 | -94.8457
2.6296 | -94.7564
2.6333 | -94.7297
2.637 | -94.7647
2.6407 | -94.8078
2.6444 | -94.8306
2.6481 | -94.8134
2.6519 | -94.7781
2.6556 | -94.7245
2.6593 | -94.6364
2.663 | -94.5178
2.6667 | -94.382
2.6704 | -94.3461
2.6741 | -94.3416
2.6778 | -94.4126
2.6815 | -94.5443
2.6852 | -94.6985
2.6889 | -94.8387
2.6926 | -94.9419
2.6963 | -95.0063
2.7 | -95.0512
2.7037 | -95.0844
2.7074 | -95.0978
2.7111 | -95.101
2.7148 | -95.1047
2.7185 | -95.1093
2.7222 | -95.1244
2.7259 | -95.1409
2.7296 | -95.1281
2.7333 | -95.0847
2.737 | -95.0202
2.7407 | -94.9642
2.7444 | -94.942
2.7481 | -94.9639
2.7519 | -95.0249
2.7556 | -95.1063
2.7593 | -95.2112
2.763 | -95.3355
2.7667 | -95.47
2.7704 | -95.5693
2.7741 | -95.5869
2.7778 | -95.5061
2.7815 | -95.34
2.7852 | -95.1177
2.7889 | -94.8791
2.7926 | -94.6702
2.7963 | -94.5324
2.8 | -94.4799
2.8037 | -94.5188
2.8074 | -94.6286
2.8111 | -94.7882
2.8148 | -94.9754
2.8185 | -95.1656
2.8222 | -95.3285
2.8259 | -95.4409
2.8296 | -95.4856
2.8333 | -95.472
2.837 | -95.4356
2.8407 | -95.374
2.8444 | -95.3103
2.8481 | -95.2618
2.8519 | -95.2253
2.8556 | -95.2203
2.8593 | -95.2641
2.863 | -95.3628
2.8667 | -95.5177
2.8704 | -95.7033
2.8741 | -95.8724
2.8778 | -95.9775
2.8815 | -96.0185
2.8852 | -95.9912
2.8889 | -95.9296
2.8926 | -95.8595
2.8963 | -95.7924
2.9 | -95.7429
2.9037 | -95.7045
2.9074 | -95.6775
2.9111 | -95.6563
2.9148 | -95.6401
2.9185 | -95.6341
2.9222 | -95.6284
2.9259 | -95.6196
2.9296 | -95.6077
2.9333 | -95.6291
2.937 | -95.7008
2.9407 | -95.8211
2.9444 | -95.9484
2.9481 | -96.0407
2.9519 | -96.0661
2.9556 | -96.0252
2.9593 | -95.9511
2.963 | -95.8822
2.9667 | -95.8398
2.9704 | -95.8272
2.9741 | -95.824
2.9778 | -95.7595
2.9815 | -95.6389
2.9852 | -95.4596
2.9889 | -95.2319
2.9926 | -95.0111
2.9963 | -94.8396
3 | -94.7701
3.0037 | -94.8254
3.0074 | -94.9876
3.0111 | -95.2071
3.0148 | -95.4369
3.0185 | -95.643
3.0222 | -95.8167
3.0259 | -96.0022
3.0296 | -96.1581
3.0333 | -96.2839
3.037 | -96.3773
3.0407 | -96.4196
3.0444 | -96.4145
3.0481 | -96.392
3.0519 | -96.3635
3.0556 | -96.3345
3.0593 | -96.3118
3.063 | -96.29
3.0667 | -96.2643
3.0704 | -96.2134
3.0741 | -96.101
3.0778 | -95.9342
3.0815 | -95.7432
3.0852 | -95.5567
3.0889 | -95.4236
3.0926 | -95.3794
3.0963 | -95.4303
3.1 | -95.5634
3.1037 | -95.7433
3.1074 | -95.9264
3.1111 | -96.0744
3.1148 | -96.1637
3.1185 | -96.1926
3.1222 | -96.1826
3.1259 | -96.1536
3.1296 | -96.1282
3.1333 | -96.1336
3.137 | -96.1586
3.1407 | -96.1839
3.1444 | -96.1972
3.1481 | -96.2004
3.1519 | -96.207
3.1556 | -96.2084
3.1593 | -96.2059
3.163 | -96.201
3.1667 | -96.1977
3.1704 | -96.1936
3.1741 | -96.1834
3.1778 | -96.1388
3.1815 | -96.0631
3.1852 | -95.9621
3.1889 | -95.8362
3.1926 | -95.7068
3.1963 | -95.6248
3.2 | -95.6255
3.2037 | -95.7849
3.2074 | -95.9121
3.2111 | -96.0722
3.2148 | -96.2393
3.2185 | -96.3918
3.2222 | -96.52
3.2259 | -96.631
3.2296 | -96.7286
3.2333 | -96.7967
3.237 | -96.8483
3.2407 | -96.8886
3.2444 | -96.9178
3.2481 | -96.9215
3.2519 | -96.9116
3.2556 | -96.9109
3.2593 | -96.9137
3.263 | -96.9234
3.2667 | -96.944
3.2704 | -96.9686
3.2741 | -96.9962
3.2778 | -97.012
3.2815 | -97.014
3.2852 | -97.0161
3.2889 | -97.0019
3.2926 | -96.9928
3.2963 | -96.9818
3.3 | -96.9567
3.3037 | -96.9243
3.3074 | -96.888
3.3111 | -96.8653
3.3148 | -96.8742
3.3185 | -96.9189
3.3222 | -96.9875
3.3259 | -97.0689
3.3296 | -97.1439
3.3333 | -97.1955
3.337 | -97.217
3.3407 | -97.1991
3.3444 | -97.1632
3.3481 | -97.0996
3.3519 | -97.0274
3.3556 | -96.9715
3.3593 | -96.948
3.363 | -96.9592
3.3667 | -96.991
3.3704 | -97.0188
3.3741 | -97.0281
3.3778 | -97.0208
3.3815 | -96.9839
3.3852 | -96.9292
3.3889 | -96.8869
3.3926 | -96.8577
3.3963 | -96.8478
3.4 | -96.8404
3.4037 | -96.8203
3.4074 | -96.7917
3.4111 | -96.7436
3.4148 | -96.7015
3.4185 | -96.6715
3.4222 | -96.6671
3.4259 | -96.7042
3.4296 | -96.7617
3.4333 | -96.8413
3.437 | -96.939
3.4407 | -97.0352
3.4444 | -97.1115
3.4481 | -97.1573
3.4519 | -97.1758
3.4556 | -97.171
3.4593 | -97.154
3.463 | -97.127
3.4667 | -97.1107
3.4704 | -97.0917
3.4741 | -97.0645
3.4778 | -97.0198
3.4815 | -96.9436
3.4852 | -96.8437
3.4889 | -96.7219
3.4926 | -96.5975
3.4963 | -96.4884
3.5 | -96.4013
3.5037 | -96.3553
3.5074 | -96.348
3.5111 | -96.3816
3.5148 | -96.4497
3.5185 | -96.5335
3.5222 | -96.6141
3.5259 | -96.6977
3.5296 | -96.7809
3.5333 | -96.8422
3.537 | -96.885
3.5407 | -96.9167
3.5444 | -96.9409
3.5481 | -96.9604
3.5519 | -96.9832
3.5556 | -97.0108
3.5593 | -97.0413
3.563 | -97.0728
3.5667 | -97.1033
3.5704 | -97.1361
3.5741 | -97.1747
3.5778 | -97.2135
3.5815 | -97.2412
3.5852 | -97.255
3.5889 | -97.254
3.5926 | -97.2478
3.5963 | -97.2399
3.6 | -97.2446
3.6037 | -97.2808
3.6074 | -97.3399
3.6111 | -97.42
3.6148 | -97.5242
3.6185 | -97.6256
3.6222 | -97.7041
3.6259 | -97.7554
3.6296 | -97.7877
3.6333 | -97.8124
3.637 | -97.8419
3.6407 | -97.8784
3.6444 | -97.9127
3.6481 | -97.932
3.6519 | -97.9426
3.6556 | -97.9228
3.6593 | -97.8773
3.663 | -97.8206
3.6667 | -97.769
3.6704 | -97.7296
3.6741 | -97.7268
3.6778 | -97.7561
3.6815 | -97.8076
3.6852 | -97.8746
3.6889 | -97.9575
3.6926 | -98.0487
3.6963 | -98.1474
3.7 | -98.2397
3.7037 | -98.3133
3.7074 | -98.3647
3.7111 | -98.3992
3.7148 | -98.421
3.7185 | -98.4137
3.7222 | -98.3737
3.7259 | -98.2929
3.7296 | -98.185
3.7333 | -98.061
3.737 | -98.0024
3.7407 | -97.953
3.7444 | -97.9508
3.7481 | -97.9856
3.7519 | -98.0357
3.7556 | -98.0902
3.7593 | -98.128
3.763 | -98.1462
3.7667 | -98.1372
3.7704 | -98.1015
3.7741 | -98.0521
3.7778 | -98.0122
3.7815 | -97.9711
3.7852 | -97.943
3.7889 | -97.9272
3.7926 | -97.9182
3.7963 | -97.9183
3.8 | -97.929
3.8037 | -97.9406
3.8074 | -97.9622
3.8111 | -97.9981
3.8148 | -98.0458
3.8185 | -98.0871
3.8222 | -98.1156
3.8259 | -98.1265
3.8296 | -98.1211
3.8333 | -98.1026
3.837 | -98.0613
3.8407 | -98.0194
3.8444 | -97.9727
3.8481 | -97.9156
3.8519 | -97.8572
3.8556 | -97.8003
3.8593 | -97.7386
3.863 | -97.6889
3.8667 | -97.6659
3.8704 | -97.6738
3.8741 | -97.7201
3.8778 | -97.8018
3.8815 | -97.9082
3.8852 | -98.0245
3.8889 | -98.1319
3.8926 | -98.2027
3.8963 | -98.241
3.9 | -98.2437
3.9037 | -98.2194
3.9074 | -98.194
3.9111 | -98.1695
3.9148 | -98.1536
3.9185 | -98.1589
3.9222 | -98.1888
3.9259 | -98.236
3.9296 | -98.2984
3.9333 | -98.3742
3.937 | -98.4608
3.9407 | -98.535
3.9444 | -98.5781
3.9481 | -98.594
3.9519 | -98.5824
3.9556 | -98.5541
3.9593 | -98.495
3.963 | -98.4117
3.9667 | -98.226
3.9704 | -97.7061
3.9741 | -97.153
3.9778 | -96.4978
3.9815 | -95.7471
3.9852 | -94.9813
3.9889 | -94.2904
3.9926 | -93.7492
3.9963 | -93.4109
4 | -93.3094
4.0037 | -93.4583
4.0074 | -93.8463
4.0111 | -94.4428
4.0148 | -95.1954
4.0185 | -96.0265
4.0222 | -96.8497
4.0259 | -97.5762
4.0296 | -98.1782
4.0333 | -98.7345
4.037 | -98.8968
4.0407 | -98.9214
4.0444 | -98.9352
4.0481 | -98.9391
4.0519 | -98.9378
4.0556 | -98.9344
4.0593 | -98.9235
4.063 | -98.9071
4.0667 | -98.878
4.0704 | -98.8337
4.0741 | -98.7666
4.0778 | -98.681
4.0815 | -98.5951
4.0852 | -98.5112
4.0889 | -98.4486
4.0926 | -98.4114
4.0963 | -98.3923
4.1 | -98.3953
4.1037 | -98.4166
4.1074 | -98.441
4.1111 | -98.471
4.1148 | -98.5099
4.1185 | -98.5571
4.1222 | -98.6204
4.1259 | -98.6887
4.1296 | -98.7613
4.1333 | -98.8285
4.137 | -98.8744
4.1407 | -98.8745
4.1444 | -98.8254
4.1481 | -98.7373
4.1519 | -98.6167
4.1556 | -98.4817
4.1593 | -98.3561
4.163 | -98.2479
4.1667 | -98.1745
4.1704 | -98.1467
4.1741 | -98.1672
4.1778 | -98.2116
4.1815 | -98.273
4.1852 | -98.3366
4.1889 | -98.3967
4.1926 | -98.4499
4.1963 | -98.4995
4.2 | -98.5472
4.2037 | -98.5984
4.2074 | -98.6422
4.2111 | -98.6745
4.2148 | -98.6872
4.2185 | -98.6696
4.2222 | -98.6169
4.2259 | -98.5428
4.2296 | -98.4461
4.2333 | -98.3379
4.237 | -98.2376
4.2407 | -98.1691
4.2444 | -98.1346
4.2481 | -98.1396
4.2519 | -98.1953
4.2556 | -98.2822
4.2593 | -98.382
4.263 | -98.4885
4.2667 | -98.5798
4.2704 | -98.6494
4.2741 | -98.673
4.2778 | -98.7142
4.2815 | -98.7484
4.2852 | -98.7837
4.2889 | -98.8194
4.2926 | -98.8509
4.2963 | -98.888
4.3 | -98.9257
4.3037 | -98.9509
4.3074 | -98.9608
4.3111 | -98.9591
4.3148 | -98.9482
4.3185 | -98.9212
4.3222 | -98.89
4.3259 | -98.8587
4.3296 | -98.8264
4.3333 | -98.8111
4.337 | -98.8163
4.3407 | -98.8473
4.3444 | -98.9042
4.3481 | -98.9744
4.3519 | -99.0535
4.3556 | -99.1298
4.3593 | -99.1901
4.363 | -99.2368
4.3667 | -99.2646
4.3704 | -99.2832
4.3741 | -99.2954
4.3778 | -99.3005
4.3815 | -99.2934
4.3852 | -99.2803
4.3889 | -99.2688
4.3926 | -99.2585
4.3963 | -99.2448
4.4 | -99.2277
4.4037 | -99.2033
4.4074 | -99.1821
4.4111 | -99.1613
4.4148 | -99.1454
4.4185 | -99.1377
4.4222 | -99.1271
4.4259 | -99.1131
4.4296 | -99.0984
4.4333 | -99.0757
4.437 | -99.0511
4.4407 | -99.0202
4.4444 | -98.9878
4.4481 | -98.9689
4.4519 | -98.9658
4.4556 | -98.9743
4.4593 | -98.9929
4.463 | -99.0117
4.4667 | -99.0371
4.4704 | -99.0583
4.4741 | -99.0743
4.4778 | -99.0807
4.4815 | -99.0849
4.4852 | -99.0826
4.4889 | -99.0788
4.4926 | -99.0783
4.4963 | -99.078
4.5 | -99.078
4.5037 | -99.0888
4.5074 | -99.1123
4.5111 | -99.1422
4.5148 | -99.1677
4.5185 | -99.1834
4.5222 | -99.1806
4.5259 | -99.1666
4.5296 | -99.1435
4.5333 | -99.133
4.537 | -99.1294
4.5407 | -99.1496
4.5444 | -99.1998
4.5481 | -99.2773
4.5519 | -99.3832
4.5556 | -99.5068
4.5593 | -99.6376
4.563 | -99.7744
4.5667 | -99.8948
4.5704 | -99.9929
4.5741 | -100.0623
4.5778 | -100.0962
4.5815 | -100.0916
4.5852 | -100.0531
4.5889 | -99.9831
4.5926 | -99.8958
4.5963 | -99.7866
4.6 | -99.6752
4.6037 | -99.5682
4.6074 | -99.4842
4.6111 | -99.4207
4.6148 | -99.3772
4.6185 | -99.3529
4.6222 | -99.3476
4.6259 | -99.3686
4.6296 | -99.4079
4.6333 | -99.4485
4.637 | -99.4985
4.6407 | -99.5551
4.6444 | -99.6052
4.6481 | -99.6612
4.6519 | -99.7189
4.6556 | -99.7738
4.6593 | -99.8175
4.663 | -99.8567
4.6667 | -99.8942
4.6704 | -99.9309
4.6741 | -99.9562
4.6778 | -99.9709
4.6815 | -99.9736
4.6852 | -99.9642
4.6889 | -99.9318
4.6926 | -99.8688
4.6963 | -99.775
4.7 | -99.6717
4.7037 | -99.5691
4.7074 | -99.4612
4.7111 | -99.3487
4.7148 | -99.245
4.7185 | -99.1514
4.7222 | -99.0868
4.7259 | -99.052
4.7296 | -99.0518
4.7333 | -99.0832
4.737 | -99.1381
4.7407 | -99.2038
4.7444 | -99.2696
4.7481 | -99.3229
4.7519 | -99.3608
4.7556 | -99.3748
4.7593 | -99.3697
4.763 | -99.3554
4.7667 | -99.3387
4.7704 | -99.3315
4.7741 | -99.3367
4.7778 | -99.3529
4.7815 | -99.3874
4.7852 | -99.4359
4.7889 | -99.496
4.7926 | -99.5621
4.7963 | -99.638
4.8 | -99.7151
4.8037 | -99.7907
4.8074 | -99.8165
4.8111 | -99.8579
4.8148 | -99.8807
4.8185 | -99.8953
4.8222 | -99.8986
4.8259 | -99.8866
4.8296 | -99.8797
4.8333 | -99.8838
4.837 | -99.8951
4.8407 | -99.908
4.8444 | -99.9168
4.8481 | -99.9305
4.8519 | -99.9505
4.8556 | -99.9549
4.8593 | -99.9492
4.863 | -99.939
4.8667 | -99.9153
4.8704 | -99.88
4.8741 | -99.8447
4.8778 | -99.8029
4.8815 | -99.7622
4.8852 | -99.7246
4.8889 | -99.6868
4.8926 | -99.6525
4.8963 | -99.623
4.9 | -99.5983
4.9037 | -99.5862
4.9074 | -99.5783
4.9111 | -99.5727
4.9148 | -99.5682
4.9185 | -99.5688
4.9222 | -99.5623
4.9259 | -99.5516
4.9296 | -99.5427
4.9333 | -99.537
4.937 | -99.5287
4.9407 | -99.5172
4.9444 | -99.5057
4.9481 | -99.4918
4.9519 | -99.4706
4.9556 | -99.4588
4.9593 | -99.3566
4.963 | -98.9247
4.9667 | -98.5248
4.9704 | -98.0623
4.9741 | -97.5012
4.9778 | -96.8782
4.9815 | -96.2418
4.9852 | -95.6469
4.9889 | -95.1408
4.9926 | -94.758
4.9963 | -94.5204
5 | -94.4443
5.0037 | -94.5335
5.0074 | -94.7832
5.0111 | -95.185
5.0148 | -95.7177
5.0185 | -96.3474
5.0222 | -97.0315
5.0259 | -97.7179
5.0296 | -98.3518
5.0333 | -98.8752
5.037 | -99.3201
5.0407 | -99.7914
5.0444 | -99.8899
5.0481 | -99.8764
5.0519 | -99.8672
5.0556 | -99.8683
5.0593 | -99.8857
5.063 | -99.9175
5.0667 | -99.9657
5.0704 | -100.0258
5.0741 | -100.0975
5.0778 | -100.1677
5.0815 | -100.2316
5.0852 | -100.2883
5.0889 | -100.3245
5.0926 | -100.3375
5.0963 | -100.3292
5.1 | -100.2949
5.1037 | -100.24
5.1074 | -100.1693
5.1111 | -100.0921
5.1148 | -100.006
5.1185 | -99.9193
5.1222 | -99.8405
5.1259 | -99.7731
5.1296 | -99.7264
5.1333 | -99.7129
5.137 | -99.725
5.1407 | -99.7632
5.1444 | -99.8224
5.1481 | -99.8959
5.1519 | -99.9816
5.1556 | -100.0721
5.1593 | -100.1558
5.163 | -100.2319
5.1667 | -100.2969
5.1704 | -100.35
5.1741 | -100.3895
5.1778 | -100.4193
5.1815 | -100.4389
5.1852 | -100.4519
5.1889 | -100.4716
5.1926 | -100.4947
5.1963 | -100.5113
5.2 | -100.5148
5.2037 | -100.5029
5.2074 | -100.4749
5.2111 | -100.4283
5.2148 | -100.3756
5.2185 | -100.3242
5.2222 | -100.2779
5.2259 | -100.2406
5.2296 | -100.2162
5.2333 | -100.205
5.237 | -100.2061
5.2407 | -100.21
5.2444 | -100.2132
5.2481 | -100.2046
5.2519 | -100.1872
5.2556 | -100.1566
5.2593 | -100.121
5.263 | -100.089
5.2667 | -100.0552
5.2704 | -100.0305
5.2741 | -100.0146
5.2778 | -100.0107
5.2815 | -100.0194
5.2852 | -100.0417
5.2889 | -100.0767
5.2926 | -100.1225
5.2963 | -100.1764
5.3 | -100.2358
5.3037 | -100.2955
5.3074 | -100.3414
5.3111 | -100.3748
5.3148 | -100.3937
5.3185 | -100.4001
5.3222 | -100.4002
5.3259 | -100.3977
5.3296 | -100.3939
5.3333 | -100.3999
5.337 | -100.4057
5.3407 | -100.4449
5.3444 | -100.4673
5.3481 | -100.4961
5.3519 | -100.5299
5.3556 | -100.5617
5.3593 | -100.5817
5.363 | -100.5913
5.3667 | -100.5787
5.3704 | -100.5578
5.3741 | -100.5309
5.3778 | -100.4972
5.3815 | -100.4577
5.3852 | -100.4252
5.3889 | -100.4098
5.3926 | -100.4101
5.3963 | -100.4249
5.4 | -100.4528
5.4037 | -100.4879
5.4074 | -100.531
5.4111 | -100.5772
5.4148 | -100.6072
5.4185 | -100.6324
5.4222 | -100.6416
5.4259 | -100.6285
5.4296 | -100.5893
5.4333 | -100.5283
5.437 | -100.457
5.4407 | -100.3764
5.4444 | -100.2954
5.4481 | -100.2182
5.4519 | -100.156
5.4556 | -100.1109
5.4593 | -100.0778
5.463 | -100.0725
5.4667 | -100.0904
5.4704 | -100.1204
5.4741 | -100.1616
5.4778 | -100.2108
5.4815 | -100.2665
5.4852 | -100.3248
5.4889 | -100.3846
5.4926 | -100.442
5.4963 | -100.4975
5.5 | -100.5571
5.5037 | -100.6014
5.5074 | -100.631
5.5111 | -100.6524
5.5148 | -100.6621
5.5185 | -100.6589
5.5222 | -100.6453
5.5259 | -100.6288
5.5296 | -100.6166
5.5333 | -100.602
5.537 | -100.5946
5.5407 | -100.5872
5.5444 | -100.5901
5.5481 | -100.5921
5.5519 | -100.5974
5.5556 | -100.6053
5.5593 | -100.6101
5.563 | -100.6175
5.5667 | -100.6269
5.5704 | -100.6307
5.5741 | -100.6293
5.5778 | -100.6183
5.5815 | -100.5919
5.5852 | -100.5527
5.5889 | -100.5025
5.5926 | -100.4483
5.5963 | -100.396
5.6 | -100.3561
5.6037 | -100.3271
5.6074 | -100.3119
5.6111 | -100.3222
5.6148 | -100.3469
5.6185 | -100.383
5.6222 | -100.422
5.6259 | -100.4626
5.6296 | -100.5011
5.6333 | -100.5293
5.637 | -100.5525
5.6407 | -100.5663
5.6444 | -100.5758
5.6481 | -100.5757
5.6519 | -100.5726
5.6556 | -100.5759
5.6593 | -100.5804
5.663 | -100.5904
5.6667 | -100.6072
5.6704 | -100.6303
5.6741 | -100.6566
5.6778 | -100.6721
5.6815 | -100.6823
5.6852 | -100.6798
5.6889 | -100.6601
5.6926 | -100.6194
5.6963 | -100.5606
5.7 | -100.4979
5.7037 | -100.4383
5.7074 | -100.3784
5.7111 | -100.3193
5.7148 | -100.2581
5.7185 | -100.2096
5.7222 | -100.1658
5.7259 | -100.1236
5.7296 | -100.0754
5.7333 | -100.0292
5.737 | -99.9802
5.7407 | -99.9285
5.7444 | -99.8769
5.7481 | -99.8235
5.7519 | -99.7754
5.7556 | -99.7261
5.7593 | -99.6769
5.763 | -99.6367
5.7667 | -99.6049
5.7704 | -99.5898
5.7741 | -99.594
5.7778 | -99.6218
5.7815 | -99.66
5.7852 | -99.701
5.7889 | -99.7467
5.7926 | -99.7955
5.7963 | -99.8373
5.8 | -99.8775
5.8037 | -99.906
5.8074 | -99.9216
5.8111 | -99.9276
5.8148 | -99.9255
5.8185 | -99.9094
5.8222 | -99.8876
5.8259 | -99.8709
5.8296 | -99.8585
5.8333 | -99.8536
5.837 | -99.8492
5.8407 | -99.8414
5.8444 | -99.8394
5.8481 | -99.8277
5.8519 | -99.8001
5.8556 | -99.7601
5.8593 | -99.7098
5.863 | -99.6513
5.8667 | -99.5905
5.8704 | -99.5291
5.8741 | -99.5038
5.8778 | -99.4757
5.8815 | -99.4584
5.8852 | -99.4573
5.8889 | -99.4563
5.8926 | -99.4592
5.8963 | -99.4703
5.9 | -99.4845
5.9037 | -99.5007
5.9074 | -99.5149
5.9111 | -99.5224
5.9148 | -99.5222
5.9185 | -99.5227
5.9222 | -99.5199
5.9259 | -99.5175
5.9296 | -99.5142
5.9333 | -99.4997
5.937 | -99.4852
5.9407 | -99.467
5.9444 | -99.447
5.9481 | -99.4219
5.9519 | -99.3858
5.9556 | -99.2978
5.9593 | -99.2298
5.963 | -99.1555
5.9667 | -99.0667
5.9704 | -98.9754
5.9741 | -98.8797
5.9778 | -98.7791
5.9815 | -98.6851
5.9852 | -98.6034
5.9889 | -98.5405
5.9926 | -98.5012
5.9963 | -98.4942
6 | -98.5242
6.0037 | -98.5919
6.0074 | -98.6932
6.0111 | -98.822
6.0148 | -98.972
6.0185 | -99.1324
6.0222 | -99.2953
6.0259 | -99.4466
6.0296 | -99.5828
6.0333 | -99.7013
6.037 | -99.7936
6.0407 | -99.8702
6.0444 | -99.9329
6.0481 | -100.012
6.0519 | -100.0246
6.0556 | -100.0225
6.0593 | -100.024
6.063 | -100.0309
6.0667 | -100.0514
6.0704 | -100.0875
6.0741 | -100.1385
6.0778 | -100.1952
6.0815 | -100.2594
6.0852 | -100.3382
6.0889 | -100.4217
6.0926 | -100.5084
6.0963 | -100.5964
6.1 | -100.6881
6.1037 | -100.7777
6.1074 | -100.8605
6.1111 | -100.9377
6.1148 | -101.014
6.1185 | -101.0817
6.1222 | -101.1279
6.1259 | -101.1621
6.1296 | -101.182
6.1333 | -101.1876
6.137 | -101.1884
6.1407 | -101.1792
6.1444 | -101.1611
6.1481 | -101.1364
6.1519 | -101.1099
6.1556 | -101.0788
6.1593 | -101.0447
6.163 | -101.0124
6.1667 | -100.985
6.1704 | -100.9617
6.1741 | -100.9365
6.1778 | -100.9088
6.1815 | -100.8833
6.1852 | -100.8675
6.1889 | -100.8579
6.1926 | -100.8431
6.1963 | -100.8378
6.2 | -100.8427
6.2037 | -100.8466
6.2074 | -100.8484
6.2111 | -100.8563
6.2148 | -100.8715
6.2185 | -100.8888
6.2222 | -100.9133
6.2259 | -100.9377
6.2296 | -100.9598
6.2333 | -100.9829
6.237 | -101.0091
6.2407 | -101.0382
6.2444 | -101.0682
6.2481 | -101.1009
6.2519 | -101.1367
6.2556 | -101.1776
6.2593 | -101.2153
6.263 | -101.2605
6.2667 | -101.3044
6.2704 | -101.3439
6.2741 | -101.3682
6.2778 | -101.3703
6.2815 | -101.3623
6.2852 | -101.3448
6.2889 | -101.3054
6.2926 | -101.2539
6.2963 | -101.1958
6.3 | -101.1319
6.3037 | -101.0605
6.3074 | -100.9837
6.3111 | -100.9108
6.3148 | -100.8413
6.3185 | -100.7804
6.3222 | -100.7312
6.3259 | -100.6944
6.3296 | -100.6709
6.3333 | -100.6558
6.337 | -100.6593
6.3407 | -100.6778
6.3444 | -100.7077
6.3481 | -100.747
6.3519 | -100.7953
6.3556 | -100.8472
6.3593 | -100.9005
6.363 | -100.9501
6.3667 | -101.005
6.3704 | -101.0541
6.3741 | -101.0944
6.3778 | -101.1284
6.3815 | -101.1512
6.3852 | -101.158
6.3889 | -101.1521
6.3926 | -101.1328
6.3963 | -101.1043
6.4 | -101.0652
6.4037 | -101.0141
6.4074 | -100.9656
6.4111 | -100.9336
6.4148 | -100.9021
6.4185 | -100.8828
6.4222 | -100.8736
6.4259 | -100.8736
6.4296 | -100.8933
6.4333 | -100.9204
6.437 | -100.9599
6.4407 | -101.009
6.4444 | -101.0665
6.4481 | -101.1274
6.4519 | -101.1891
6.4556 | -101.2548
6.4593 | -101.319
6.463 | -101.3691
6.4667 | -101.4078
6.4704 | -101.4415
6.4741 | -101.4671
6.4778 | -101.4868
6.4815 | -101.5017
6.4852 | -101.5079
6.4889 | -101.5055
6.4926 | -101.4951
6.4963 | -101.4821
6.5 | -101.4682
6.5037 | -101.4474
6.5074 | -101.4261
6.5111 | -101.4047
6.5148 | -101.3831
6.5185 | -101.3583
6.5222 | -101.3392
6.5259 | -101.3193
6.5296 | -101.2945
6.5333 | -101.2764
6.537 | -101.2575
6.5407 | -101.2371
6.5444 | -101.2215
6.5481 | -101.2108
6.5519 | -101.1984
6.5556 | -101.1944
6.5593 | -101.197
6.563 | -101.21
6.5667 | -101.2349
6.5704 | -101.2637
6.5741 | -101.2961
6.5778 | -101.3332
6.5815 | -101.3802
6.5852 | -101.4315
6.5889 | -101.4842
6.5926 | -101.5407
6.5963 | -101.5937
6.6 | -101.6475
6.6037 | -101.7008
6.6074 | -101.7536
6.6111 | -101.8011
6.6148 | -101.8357
6.6185 | -101.8594
6.6222 | -101.8756
6.6259 | -101.883
6.6296 | -101.8828
6.6333 | -101.8694
6.637 | -101.8457
6.6407 | -101.8239
6.6444 | -101.7993
6.6481 | -101.7693
6.6519 | -101.7458
6.6556 | -101.7259
6.6593 | -101.7179
6.663 | -101.7201
6.6667 | -101.7339
6.6704 | -101.7556
6.6741 | -101.789
6.6778 | -101.8242
6.6815 | -101.8575
6.6852 | -101.8936
6.6889 | -101.9272
6.6926 | -101.9517
6.6963 | -101.9726
6.7 | -101.985
6.7037 | -101.9927
6.7074 | -101.9896
6.7111 | -101.9857
6.7148 | -101.9853
6.7185 | -101.9817
6.7222 | -101.9686
6.7259 | -101.9545
6.7296 | -101.94
6.7333 | -101.9308
6.737 | -101.9303
6.7407 | -101.9271
6.7444 | -101.9297
6.7481 | -101.9459
6.7519 | -101.9602
6.7556 | -101.972
6.7593 | -101.9857
6.763 | -101.9999
6.7667 | -102.0124
6.7704 | -102.0302
6.7741 | -102.0551
6.7778 | -102.0862
6.7815 | -102.1131
6.7852 | -102.1447
6.7889 | -102.1795
6.7926 | -102.2094
6.7963 | -102.2364
6.8 | -102.2565
6.8037 | -102.2709
6.8074 | -102.2881
6.8111 | -102.2908
6.8148 | -102.2799
6.8185 | -102.2631
6.8222 | -102.2411
6.8259 | -102.2128
6.8296 | -102.1802
6.8333 | -102.14
6.837 | -102.0997
6.8407 | -102.0552
6.8444 | -102.0033
6.8481 | -101.9524
6.8519 | -101.9149
6.8556 | -101.8859
6.8593 | -101.8625
6.863 | -101.8517
6.8667 | -101.8529
6.8704 | -101.8632
6.8741 | -101.8831
6.8778 | -101.9124
6.8815 | -101.9465
6.8852 | -101.979
6.8889 | -102.016
6.8926 | -102.0577
6.8963 | -102.0881
6.9 | -102.1079
6.9037 | -102.1181
6.9074 | -102.1159
6.9111 | -102.1057
6.9148 | -102.0886
6.9185 | -102.0596
6.9222 | -102.0227
6.9259 | -101.9748
6.9296 | -101.9145
6.9333 | -101.854
6.937 | -101.794
6.9407 | -101.7391
6.9444 | -101.7139
6.9481 | -101.6729
6.9519 | -101.6438
6.9556 | -101.6221
6.9593 | -101.6104
6.963 | -101.6068
6.9667 | -101.6064
6.9704 | -101.6067
6.9741 | -101.6092
6.9778 | -101.6153
6.9815 | -101.6254
6.9852 | -101.6395
6.9889 | -101.6546
6.9926 | -101.6729
6.9963 | -101.702
7 | -101.7441
7.0037 | -101.7986
7.0074 | -101.8591
7.0111 | -101.9228
7.0148 | -101.9886
7.0185 | -102.0526
7.0222 | -102.1176
7.0259 | -102.176
7.0296 | -102.2278
7.0333 | -102.2796
7.037 | -102.3249
7.0407 | -102.3595
7.0444 | -102.387
7.0481 | -102.4065
7.0519 | -102.4269
7.0556 | -102.4465
7.0593 | -102.4519
7.063 | -102.4476
7.0667 | -102.4384
7.0704 | -102.4256
7.0741 | -102.4109
7.0778 | -102.3939
7.0815 | -102.3805
7.0852 | -102.3659
7.0889 | -102.3464
7.0926 | -102.3267
7.0963 | -102.313
7.1 | -102.2942
7.1037 | -102.2714
7.1074 | -102.2452
7.1111 | -102.2037
7.1148 | -102.1589
7.1185 | -102.1151
7.1222 | -102.0634
7.1259 | -102.0092
7.1296 | -101.9517
7.1333 | -101.9006
7.137 | -101.8581
7.1407 | -101.819
7.1444 | -101.7865
7.1481 | -101.768
7.1519 | -101.7635
7.1556 | -101.7663
7.1593 | -101.778
7.163 | -101.8023
7.1667 | -101.8344
7.1704 | -101.8827
7.1741 | -101.9407
7.1778 | -102.0042
7.1815 | -102.0697
7.1852 | -102.1381
7.1889 | -102.1998
7.1926 | -102.2579
7.1963 | -102.3141
7.2 | -102.3633
7.2037 | -102.3935
7.2074 | -102.4179
7.2111 | -102.4338
7.2148 | -102.4447
7.2185 | -102.4561
7.2222 | -102.4568
7.2259 | -102.4495
7.2296 | -102.4436
7.2333 | -102.4325
7.237 | -102.4185
7.2407 | -102.3967
7.2444 | -102.3804
7.2481 | -102.3697
7.2519 | -102.3594
7.2556 | -102.3512
7.2593 | -102.3444
7.263 | -102.3389
7.2667 | -102.341
7.2704 | -102.3432
7.2741 | -102.3439
7.2778 | -102.3405
7.2815 | -102.3367
7.2852 | -102.3314
7.2889 | -102.3223
7.2926 | -102.305
7.2963 | -102.2891
7.3 | -102.2714
7.3037 | -102.2451
7.3074 | -102.2137
7.3111 | -102.1845
7.3148 | -102.1509
7.3185 | -102.1122
7.3222 | -102.0778
7.3259 | -102.0442
7.3296 | -102.0109
7.3333 | -101.9829
7.337 | -101.9581
7.3407 | -101.9433
7.3444 | -101.9354
7.3481 | -101.9373
7.3519 | -101.9433
7.3556 | -101.9539
7.3593 | -101.9743
7.363 | -102.003
7.3667 | -102.0439
7.3704 | -102.0903
7.3741 | -102.1316
7.3778 | -102.1682
7.3815 | -102.2066
7.3852 | -102.2417
7.3889 | -102.2726
7.3926 | -102.3027
7.3963 | -102.3281
7.4 | -102.3511
7.4037 | -102.3644
7.4074 | -102.3759
7.4111 | -102.3883
7.4148 | -102.3957
7.4185 | -102.3952
7.4222 | -102.3923
7.4259 | -102.3919
7.4296 | -102.3888
7.4333 | -102.3915
7.437 | -102.4004
7.4407 | -102.4089
7.4444 | -102.4272
7.4481 | -102.4487
7.4519 | -102.4805
7.4556 | -102.5189
7.4593 | -102.5511
7.463 | -102.5895
7.4667 | -102.6279
7.4704 | -102.6586
7.4741 | -102.6773
7.4778 | -102.6628
7.4815 | -102.6541
7.4852 | -102.6428
7.4889 | -102.6308
7.4926 | -102.6074
7.4963 | -102.5767
7.5 | -102.5421
7.5037 | -102.505
7.5074 | -102.4749
7.5111 | -102.447
7.5148 | -102.4274
7.5185 | -102.417
7.5222 | -102.4202
7.5259 | -102.4366
7.5296 | -102.467
7.5333 | -102.5084
7.537 | -102.5592
7.5407 | -102.6188
7.5444 | -102.6848
7.5481 | -102.7541
7.5519 | -102.8275
7.5556 | -102.9055
7.5593 | -102.9855
7.563 | -103.0605
7.5667 | -103.1358
7.5704 | -103.2189
7.5741 | -103.2938
7.5778 | -103.3612
7.5815 | -103.4194
7.5852 | -103.4675
7.5889 | -103.5055
7.5926 | -103.5337
7.5963 | -103.5477
7.6 | -103.5418
7.6037 | -103.5155
7.6074 | -103.4717
7.6111 | -103.4176
7.6148 | -103.3577
7.6185 | -103.2915
7.6222 | -103.2138
7.6259 | -103.1349
7.6296 | -103.0642
7.6333 | -102.9926
7.637 | -102.9253
7.6407 | -102.8586
7.6444 | -102.794
7.6481 | -102.7371
7.6519 | -102.689
7.6556 | -102.6466
7.6593 | -102.6077
7.663 | -102.5771
7.6667 | -102.5545
7.6704 | -102.532
7.6741 | -102.5082
7.6778 | -102.4878
7.6815 | -102.4685
7.6852 | -102.4451
7.6889 | -102.4225
7.6926 | -102.4021
7.6963 | -102.3861
7.7 | -102.3715
7.7037 | -102.3582
7.7074 | -102.3491
7.7111 | -102.3493
7.7148 | -102.3557
7.7185 | -102.3693
7.7222 | -102.392
7.7259 | -102.4179
7.7296 | -102.4462
7.7333 | -102.4784
7.737 | -102.513
7.7407 | -102.5463
7.7444 | -102.5721
7.7481 | -102.599
7.7519 | -102.6218
7.7556 | -102.6462
7.7593 | -102.6691
7.763 | -102.6837
7.7667 | -102.6897
7.7704 | -102.6945
7.7741 | -102.6959
7.7778 | -102.6919
7.7815 | -102.6861
7.7852 | -102.6822
7.7889 | -102.6826
7.7926 | -102.683
7.7963 | -102.686
7.8 | -102.6964
7.8037 | -102.7086
7.8074 | -102.7185
7.8111 | -102.7295
7.8148 | -102.7424
7.8185 | -102.7511
7.8222 | -102.7604
7.8259 | -102.7677
7.8296 | -102.776
7.8333 | -102.7841
7.837 | -102.7924
7.8407 | -102.8016
7.8444 | -102.816
7.8481 | -102.835
7.8519 | -102.8574
7.8556 | -102.8784
7.8593 | -102.8975
7.863 | -102.914
7.8667 | -102.929
7.8704 | -102.9435
7.8741 | -102.9574
7.8778 | -102.9684
7.8815 | -102.9787
7.8852 | -102.9803
7.8889 | -102.9817
7.8926 | -102.9856
7.8963 | -102.9876
7.9 | -102.9882
7.9037 | -102.9936
7.9074 | -103.0028
7.9111 | -103.0134
7.9148 | -103.0261
7.9185 | -103.0433
7.9222 | -103.0731
7.9259 | -103.1026
7.9296 | -103.1322
7.9333 | -103.1656
7.937 | -103.1952
7.9407 | -103.2004
7.9444 | -103.1379
7.9481 | -103.0903
7.9519 | -103.0354
7.9556 | -102.9572
7.9593 | -102.8467
7.963 | -102.7135
7.9667 | -102.5622
7.9704 | -102.398
7.9741 | -102.2237
7.9778 | -102.0471
7.9815 | -101.879
7.9852 | -101.7241
7.9889 | -101.5909
7.9926 | -101.4857
7.9963 | -101.4112
8 | -101.373
8.0037 | -101.3667
8.0074 | -101.3874
8.0111 | -101.4395
8.0148 | -101.571
8.0185 | -101.6577
8.0222 | -101.7545
8.0259 | -101.8617
8.0296 | -101.9765
8.0333 | -102.0891
8.037 | -102.1969
8.0407 | -102.2955
8.0444 | -102.3844
8.0481 | -102.4622
8.0519 | -102.5274
8.0556 | -102.5824
8.0593 | -102.636
8.063 | -102.709
8.0667 | -102.731
8.0704 | -102.7409
8.0741 | -102.7524
8.0778 | -102.7699
8.0815 | -102.7847
8.0852 | -102.8004
8.0889 | -102.8179
8.0926 | -102.8413
8.0963 | -102.8668
8.1 | -102.8879
8.1037 | -102.9131
8.1074 | -102.937
8.1111 | -102.9546
8.1148 | -102.9664
8.1185 | -102.9785
8.1222 | -102.984
8.1259 | -102.9892
8.1296 | -102.99
8.1333 | -102.9855
8.137 | -102.9737
8.1407 | -102.9537
8.1444 | -102.9389
8.1481 | -102.9168
8.1519 | -102.8881
8.1556 | -102.864
8.1593 | -102.8401
8.163 | -102.813
8.1667 | -102.7896
8.1704 | -102.7713
8.1741 | -102.7582
8.1778 | -102.7463
8.1815 | -102.7387
8.1852 | -102.7294
8.1889 | -102.7251
8.1926 | -102.7218
8.1963 | -102.7205
8.2 | -102.7231
8.2037 | -102.7316
8.2074 | -102.7409
8.2111 | -102.7521
8.2148 | -102.7671
8.2185 | -102.7862
8.2222 | -102.8092
8.2259 | -102.8343
8.2296 | -102.8619
8.2333 | -102.8908
8.237 | -102.9198
8.2407 | -102.9419
8.2444 | -102.9643
8.2481 | -102.9859
8.2519 | -103.0037
8.2556 | -103.0208
8.2593 | -103.0364
8.263 | -103.0487
8.2667 | -103.0529
8.2704 | -103.0581
8.2741 | -103.0613
8.2778 | -103.0567
8.2815 | -103.0512
8.2852 | -103.047
8.2889 | -103.0407
8.2926 | -103.0276
8.2963 | -103.0145
8.3 | -102.9999
8.3037 | -102.9909
8.3074 | -102.9833
8.3111 | -102.9718
8.3148 | -102.9565
8.3185 | -102.9493
8.3222 | -102.9382
8.3259 | -102.9264
8.3296 | -102.9206
8.3333 | -102.9114
8.337 | -102.9035
8.3407 | -102.8984
8.3444 | -102.8911
8.3481 | -102.8846
8.3519 | -102.8825
8.3556 | -102.8831
8.3593 | -102.8785
8.363 | -102.8681
8.3667 | -102.856
8.3704 | -102.8373
8.3741 | -102.8191
8.3778 | -102.7963
8.3815 | -102.774
8.3852 | -102.7553
8.3889 | -102.7364
8.3926 | -102.7177
8.3963 | -102.7011
8.4 | -102.6908
8.4037 | -102.681
8.4074 | -102.6749
8.4111 | -102.6735
8.4148 | -102.6771
8.4185 | -102.6861
8.4222 | -102.7008
8.4259 | -102.7173
8.4296 | -102.7437
8.4333 | -102.7718
8.437 | -102.8061
8.4407 | -102.84
8.4444 | -102.8794
8.4481 | -102.9212
8.4519 | -102.9574
8.4556 | -102.9906
8.4593 | -103.0195
8.463 | -103.0376
8.4667 | -103.0544
8.4704 | -103.0619
8.4741 | -103.0612
8.4778 | -103.0591
8.4815 | -103.0545
8.4852 | -103.0472
8.4889 | -103.0371
8.4926 | -103.0252
8.4963 | -103.0114
8.5 | -103.0015
8.5037 | -102.99
8.5074 | -102.9762
8.5111 | -102.9641
8.5148 | -102.9558
8.5185 | -102.9477
8.5222 | -102.9411
8.5259 | -102.9436
8.5296 | -102.9505
8.5333 | -102.9613
8.537 | -102.9772
8.5407 | -102.9952
8.5444 | -103.0249
8.5481 | -103.0647
8.5519 | -103.1006
8.5556 | -103.1363
8.5593 | -103.1701
8.563 | -103.206
8.5667 | -103.2385
8.5704 | -103.2677
8.5741 | -103.2911
8.5778 | -103.3056
8.5815 | -103.3141
8.5852 | -103.3128
8.5889 | -103.3058
8.5926 | -103.2902
8.5963 | -103.2664
8.6 | -103.2384
8.6037 | -103.2086
8.6074 | -103.1764
8.6111 | -103.1439
8.6148 | -103.1094
8.6185 | -103.0803
8.6222 | -103.0571
8.6259 | -103.037
8.6296 | -103.0203
8.6333 | -103.0079
8.637 | -102.9971
8.6407 | -102.9875
8.6444 | -102.9818
8.6481 | -102.984
8.6519 | -102.9867
8.6556 | -102.993
8.6593 | -103.0086
8.663 | -103.0334
8.6667 | -103.0598
8.6704 | -103.0848
8.6741 | -103.1168
8.6778 | -103.145
8.6815 | -103.1679
8.6852 | -103.1907
8.6889 | -103.2173
8.6926 | -103.2369
8.6963 | -103.2512
8.7 | -103.2606
8.7037 | -103.269
8.7074 | -103.2758
8.7111 | -103.2785
8.7148 | -103.2763
8.7185 | -103.2768
8.7222 | -103.2774
8.7259 | -103.2762
8.7296 | -103.2715
8.7333 | -103.265
8.737 | -103.2602
8.7407 | -103.254
8.7444 | -103.2476
8.7481 | -103.2401
8.7519 | -103.2268
8.7556 | -103.2144
8.7593 | -103.2073
8.763 | -103.201
8.7667 | -103.1999
8.7704 | -103.197
8.7741 | -103.1952
8.7778 | -103.1873
8.7815 | -103.1759
8.7852 | -103.1676
8.7889 | -103.1553
8.7926 | -103.137
8.7963 | -103.1215
8.8 | -103.1061
8.8037 | -103.0877
8.8074 | -103.0734
8.8111 | -103.0553
8.8148 | -103.0349
8.8185 | -103.0171
8.8222 | -103.0006
8.8259 | -102.9883
8.8296 | -102.9799
8.8333 | -102.9775
8.837 | -102.98
8.8407 | -102.9893
8.8444 | -103.0032
8.8481 | -103.0211
8.8519 | -103.0472
8.8556 | -103.0762
8.8593 | -103.1068
8.863 | -103.1415
8.8667 | -103.1787
8.8704 | -103.2164
8.8741 | -103.2495
8.8778 | -103.2753
8.8815 | -103.2961
8.8852 | -103.3126
8.8889 | -103.3148
8.8926 | -103.3068
8.8963 | -103.2871
8.9 | -103.2615
8.9037 | -103.2311
8.9074 | -103.1955
8.9111 | -103.1599
8.9148 | -103.1179
8.9185 | -103.0705
8.9222 | -103.0257
8.9259 | -102.9803
8.9296 | -102.9358
8.9333 | -102.8899
8.937 | -102.8312
8.9407 | -102.785
8.9444 | -102.7403
8.9481 | -102.6971
8.9519 | -102.653
8.9556 | -102.612
8.9593 | -102.5725
8.963 | -102.5301
8.9667 | -102.4904
8.9704 | -102.4523
8.9741 | -102.417
8.9778 | -102.3835
8.9815 | -102.3518
8.9852 | -102.3316
8.9889 | -102.3185
8.9926 | -102.3135
8.9963 | -102.3182
9 | -102.3346
9.0037 | -102.3631
9.0074 | -102.4047
9.0111 | -102.4581
9.0148 | -102.5155
9.0185 | -102.5835
9.0222 | -102.6554
9.0259 | -102.732
9.0296 | -102.806
9.0333 | -102.875
9.037 | -102.9425
9.0407 | -103.0044
9.0444 | -103.0555
9.0481 | -103.0986
9.0519 | -103.1352
9.0556 | -103.162
9.0593 | -103.1826
9.063 | -103.2014
9.0667 | -103.2239
9.0704 | -103.2273
9.0741 | -103.2269
9.0778 | -103.2241
9.0815 | -103.2281
9.0852 | -103.2274
9.0889 | -103.2277
9.0926 | -103.2287
9.0963 | -103.2286
9.1 | -103.2303
9.1037 | -103.2289
9.1074 | -103.2259
9.1111 | -103.2295
9.1148 | -103.2311
9.1185 | -103.2229
9.1222 | -103.2162
9.1259 | -103.2039
9.1296 | -103.1928
9.1333 | -103.179
9.137 | -103.1647
9.1407 | -103.1494
9.1444 | -103.135
9.1481 | -103.1256
9.1519 | -103.1186
9.1556 | -103.1168
9.1593 | -103.115
9.163 | -103.119
9.1667 | -103.1255
9.1704 | -103.1318
9.1741 | -103.1443
9.1778 | -103.1631
9.1815 | -103.1877
9.1852 | -103.2105
9.1889 | -103.2361
9.1926 | -103.2658
9.1963 | -103.2979
9.2 | -103.3309
9.2037 | -103.357
9.2074 | -103.3792
9.2111 | -103.3997
9.2148 | -103.4185
9.2185 | -103.4344
9.2222 | -103.4453
9.2259 | -103.4462
9.2296 | -103.4432
9.2333 | -103.4405
9.237 | -103.4312
9.2407 | -103.4206
9.2444 | -103.4142
9.2481 | -103.4017
9.2519 | -103.392
9.2556 | -103.3923
9.2593 | -103.392
9.263 | -103.3975
9.2667 | -103.4052
9.2704 | -103.4197
9.2741 | -103.4394
9.2778 | -103.4621
9.2815 | -103.4897
9.2852 | -103.5212
9.2889 | -103.554
9.2926 | -103.5852
9.2963 | -103.6165
9.3 | -103.6478
9.3037 | -103.6785
9.3074 | -103.7062
9.3111 | -103.7247
9.3148 | -103.742
9.3185 | -103.7548
9.3222 | -103.7585
9.3259 | -103.76
9.3296 | -103.7532
9.3333 | -103.7413
9.337 | -103.7258
9.3407 | -103.7144
9.3444 | -103.6993
9.3481 | -103.6758
9.3519 | -103.6453
9.3556 | -103.613
9.3593 | -103.586
9.363 | -103.5615
9.3667 | -103.5407
9.3704 | -103.5198
9.3741 | -103.5067
9.3778 | -103.5004
9.3815 | -103.4927
9.3852 | -103.4897
9.3889 | -103.492
9.3926 | -103.4979
9.3963 | -103.5056
9.4 | -103.5141
9.4037 | -103.5234
9.4074 | -103.5287
9.4111 | -103.5323
9.4148 | -103.5329
9.4185 | -103.5246
9.4222 | -103.5093
9.4259 | -103.4873
9.4296 | -103.4646
9.4333 | -103.4362
9.437 | -103.3997
9.4407 | -103.3602
9.4444 | -103.3156
9.4481 | -103.2674
9.4519 | -103.2179
9.4556 | -103.1645
9.4593 | -103.1177
9.463 | -103.0728
9.4667 | -103.0336
9.4704 | -103.0014
9.4741 | -102.9773
9.4778 | -102.9562
9.4815 | -102.9423
9.4852 | -102.9379
9.4889 | -102.9458
9.4926 | -102.9585
9.4963 | -102.971
9.5 | -102.9888
9.5037 | -103.0041
9.5074 | -103.0228
9.5111 | -103.0388
9.5148 | -103.0497
9.5185 | -103.0614
9.5222 | -103.0717
9.5259 | -103.0758
9.5296 | -103.0712
9.5333 | -103.0646
9.537 | -103.0568
9.5407 | -103.0436
9.5444 | -103.0318
9.5481 | -103.0183
9.5519 | -103.0018
9.5556 | -102.9899
9.5593 | -102.9823
9.563 | -102.9793
9.5667 | -102.9781
9.5704 | -102.9774
9.5741 | -102.9817
9.5778 | -102.9885
9.5815 | -103.0007
9.5852 | -103.0209
9.5889 | -103.0447
9.5926 | -103.074
9.5963 | -103.1043
9.6 | -103.1384
9.6037 | -103.1768
9.6074 | -103.2183
9.6111 | -103.2602
9.6148 | -103.3012
9.6185 | -103.344
9.6222 | -103.3838
9.6259 | -103.4227
9.6296 | -103.4607
9.6333 | -103.4999
9.637 | -103.5389
9.6407 | -103.5739
9.6444 | -103.6085
9.6481 | -103.6381
9.6519 | -103.6685
9.6556 | -103.7008
9.6593 | -103.7313
9.663 | -103.761
9.6667 | -103.7898
9.6704 | -103.8201
9.6741 | -103.8439
9.6778 | -103.862
9.6815 | -103.8783
9.6852 | -103.8906
9.6889 | -103.9005
9.6926 | -103.9119
9.6963 | -103.9132
9.7 | -103.9098
9.7037 | -103.906
9.7074 | -103.9002
9.7111 | -103.8975
9.7148 | -103.8857
9.7185 | -103.8681
9.7222 | -103.8511
9.7259 | -103.8277
9.7296 | -103.7951
9.7333 | -103.7643
9.737 | -103.7329
9.7407 | -103.6975
9.7444 | -103.6645
9.7481 | -103.6331
9.7519 | -103.605
9.7556 | -103.5743
9.7593 | -103.5414
9.763 | -103.5111
9.7667 | -103.4895
9.7704 | -103.4695
9.7741 | -103.45
9.7778 | -103.4351
9.7815 | -103.4271
9.7852 | -103.4248
9.7889 | -103.4262
9.7926 | -103.4345
9.7963 | -103.4458
9.8 | -103.4609
9.8037 | -103.4822
9.8074 | -103.5113
9.8111 | -103.5468
9.8148 | -103.5854
9.8185 | -103.6271
9.8222 | -103.6721
9.8259 | -103.7233
9.8296 | -103.7734
9.8333 | -103.8222
9.837 | -103.8664
9.8407 | -103.9109
9.8444 | -103.9537
9.8481 | -103.9881
9.8519 | -104.0149
9.8556 | -104.0403
9.8593 | -104.0669
9.863 | -104.0882
9.8667 | -104.0975
9.8704 | -104.1001
9.8741 | -104.0994
9.8778 | -104.0886
9.8815 | -104.0731
9.8852 | -104.0526
9.8889 | -104.0251
9.8926 | -103.9954
9.8963 | -103.9693
9.9 | -103.9472
9.9037 | -103.9286
9.9074 | -103.9121
9.9111 | -103.8893
9.9148 | -103.8719
9.9185 | -103.8572
9.9222 | -103.8455
9.9259 | -103.824
9.9296 | -103.7745
9.9333 | -103.7451
9.937 | -103.7154
9.9407 | -103.6822
9.9444 | -103.6444
9.9481 | -103.5943
9.9519 | -103.5334
9.9556 | -103.4662
9.9593 | -103.3926
9.963 | -103.3139
9.9667 | -103.2311
9.9704 | -103.1441
9.9741 | -103.0576
9.9778 | -102.9733
9.9815 | -102.8978
9.9852 | -102.8285
9.9889 | -102.769
9.9926 | -102.7252
9.9963 | -102.6972
10 | -102.6788
10.0037 | -102.6735
10.0074 | -102.6839
10.0111 | -102.7137
10.0148 | -102.7572
10.0185 | -102.8135
10.0222 | -102.8786
10.0259 | -102.9507
10.0296 | -103.0287
10.0333 | -103.1132
10.037 | -103.2022
10.0407 | -103.2882
10.0444 | -103.3721
10.0481 | -103.4543
10.0519 | -103.5323
10.0556 | -103.6109
10.0593 | -103.6788
10.063 | -103.7411
10.0667 | -103.7977
10.0704 | -103.8572
10.0741 | -103.9309
10.0778 | -103.9702
10.0815 | -103.9982
10.0852 | -104.0218
10.0889 | -104.0342
10.0926 | -104.0471
10.0963 | -104.0595
10.1 | -104.0669
10.1037 | -104.0742
10.1074 | -104.0781
10.1111 | -104.0706
10.1148 | -104.0613
10.1185 | -104.0509
10.1222 | -104.0389
10.1259 | -104.0216
10.1296 | -103.9985
10.1333 | -103.9737
10.137 | -103.9466
10.1407 | -103.9175
10.1444 | -103.8916
10.1481 | -103.8666
10.1519 | -103.8445
10.1556 | -103.824
10.1593 | -103.8029
10.163 | -103.7838
10.1667 | -103.766
10.1704 | -103.7518
10.1741 | -103.7414
10.1778 | -103.7367
10.1815 | -103.7296
10.1852 | -103.7235
10.1889 | -103.7223
10.1926 | -103.7239
10.1963 | -103.7234
10.2 | -103.724
10.2037 | -103.7231
10.2074 | -103.7266
10.2111 | -103.7325
10.2148 | -103.7335
10.2185 | -103.7343
10.2222 | -103.7323
10.2259 | -103.7298
10.2296 | -103.7273
10.2333 | -103.7257
10.237 | -103.7295
10.2407 | -103.7305
10.2444 | -103.7281
10.2481 | -103.7229
10.2519 | -103.7164
10.2556 | -103.712
10.2593 | -103.7128
10.263 | -103.7122
10.2667 | -103.7119
10.2704 | -103.7111
10.2741 | -103.7105
10.2778 | -103.7131
10.2815 | -103.7206
10.2852 | -103.7296
10.2889 | -103.7406
10.2926 | -103.7545
10.2963 | -103.7712
10.3 | -103.7917
10.3037 | -103.8134
10.3074 | -103.8286
10.3111 | -103.8419
10.3148 | -103.8581
10.3185 | -103.8744
10.3222 | -103.8859
10.3259 | -103.894
10.3296 | -103.8986
10.3333 | -103.9014
10.337 | -103.8982
10.3407 | -103.8938
10.3444 | -103.8875
10.3481 | -103.8797
10.3519 | -103.8665
10.3556 | -103.8546
10.3593 | -103.8401
10.363 | -103.8229
10.3667 | -103.8064
10.3704 | -103.786
10.3741 | -103.7692
10.3778 | -103.7523
10.3815 | -103.738
10.3852 | -103.7246
10.3889 | -103.7149
10.3926 | -103.7089
10.3963 | -103.7062
10.4 | -103.7101
10.4037 | -103.7151
10.4074 | -103.7208
10.4111 | -103.7284
10.4148 | -103.7361
10.4185 | -103.7473
10.4222 | -103.7598
10.4259 | -103.7746
10.4296 | -103.7921
10.4333 | -103.8013
10.437 | -103.8095
10.4407 | -103.8216
10.4444 | -103.832
10.4481 | -103.8361
10.4519 | -103.8383
10.4556 | -103.8451
10.4593 | -103.8555
10.463 | -103.8661
10.4667 | -103.8768
10.4704 | -103.887
10.4741 | -103.8965
10.4778 | -103.9043
10.4815 | -103.9142
10.4852 | -103.9253
10.4889 | -103.9401
10.4926 | -103.9518
10.4963 | -103.9641
10.5 | -103.9771
10.5037 | -103.9926
10.5074 | -104.0037
10.5111 | -104.0161
10.5148 | -104.0282
10.5185 | -104.0395
10.5222 | -104.0488
10.5259 | -104.0556
10.5296 | -104.0584
10.5333 | -104.0536
10.537 | -104.05
10.5407 | -104.0407
10.5444 | -104.029
10.5481 | -104.0162
10.5519 | -104.0018
10.5556 | -103.9839
10.5593 | -103.9661
10.563 | -103.9473
10.5667 | -103.9292
10.5704 | -103.9091
10.5741 | -103.8909
10.5778 | -103.8745
10.5815 | -103.8562
10.5852 | -103.8406
10.5889 | -103.8307
10.5926 | -103.8224
10.5963 | -103.8172
10.6 | -103.8188
10.6037 | -103.8227
10.6074 | -103.8278
10.6111 | -103.8345
10.6148 | -103.8373
10.6185 | -103.8383
10.6222 | -103.8441
10.6259 | -103.8552
10.6296 | -103.8669
10.6333 | -103.8765
10.637 | -103.8885
10.6407 | -103.8964
10.6444 | -103.9052
10.6481 | -103.9106
10.6519 | -103.9159
10.6556 | -103.9185
10.6593 | -103.921
10.663 | -103.9211
10.6667 | -103.9246
10.6704 | -103.9286
10.6741 | -103.9304
10.6778 | -103.9326
10.6815 | -103.9402
10.6852 | -103.9514
10.6889 | -103.9601
10.6926 | -103.9664
10.6963 | -103.975
10.7 | -103.9817
10.7037 | -103.9871
10.7074 | -103.9911
10.7111 | -103.9973
10.7148 | -104.0032
10.7185 | -104.0034
10.7222 | -104.0041
10.7259 | -104.0034
10.7296 | -103.9995
10.7333 | -103.9945
10.737 | -103.9928
10.7407 | -103.9884
10.7444 | -103.9803
10.7481 | -103.9714
10.7519 | -103.9634
10.7556 | -103.9588
10.7593 | -103.9532
10.763 | -103.9506
10.7667 | -103.9514
10.7704 | -103.9588
10.7741 | -103.9649
10.7778 | -103.9701
10.7815 | -103.9749
10.7852 | -103.9838
10.7889 | -103.9975
10.7926 | -104.0112
10.7963 | -104.0297
10.8 | -104.049
10.8037 | -104.0733
10.8074 | -104.0986
10.8111 | -104.1247
10.8148 | -104.1499
10.8185 | -104.1773
10.8222 | -104.2018
10.8259 | -104.2248
10.8296 | -104.2509
10.8333 | -104.2714
10.837 | -104.2909
10.8407 | -104.3061
10.8444 | -104.3217
10.8481 | -104.3357
10.8519 | -104.3499
10.8556 | -104.359
10.8593 | -104.3633
10.863 | -104.3676
10.8667 | -104.3663
10.8704 | -104.3636
10.8741 | -104.3589
10.8778 | -104.3527
10.8815 | -104.3427
10.8852 | -104.3313
10.8889 | -104.3204
10.8926 | -104.3095
10.8963 | -104.2945
10.9 | -104.2767
10.9037 | -104.2626
10.9074 | -104.2444
10.9111 | -104.2239
10.9148 | -104.2029
10.9185 | -104.1745
10.9222 | -104.1218
10.9259 | -104.0834
10.9296 | -104.0443
10.9333 | -104.0042
10.937 | -103.9629
10.9407 | -103.922
10.9444 | -103.8785
10.9481 | -103.8323
10.9519 | -103.7846
10.9556 | -103.7384
10.9593 | -103.6919
10.963 | -103.647
10.9667 | -103.6057
10.9704 | -103.5668
10.9741 | -103.5337
10.9778 | -103.5009
10.9815 | -103.4733
10.9852 | -103.4553
10.9889 | -103.4427
10.9926 | -103.4375
10.9963 | -103.4398
11 | -103.4514
11.0037 | -103.4707
11.0074 | -103.4985
11.0111 | -103.5308
11.0148 | -103.5727
11.0185 | -103.6251
11.0222 | -103.6802
11.0259 | -103.7411
11.0296 | -103.8038
11.0333 | -103.8661
11.037 | -103.931
11.0407 | -103.993
11.0444 | -104.0496
11.0481 | -104.1056
11.0519 | -104.1566
11.0556 | -104.2025
11.0593 | -104.2455
11.063 | -104.2722
11.0667 | -104.294
11.0704 | -104.3083
11.0741 | -104.3117
11.0778 | -104.3176
11.0815 | -104.336
11.0852 | -104.3289
11.0889 | -104.3176
11.0926 | -104.3053
11.0963 | -104.2906
11.1 | -104.2705
11.1037 | -104.2539
11.1074 | -104.2437
11.1111 | -104.2385
11.1148 | -104.2353
11.1185 | -104.2385
11.1222 | -104.239
11.1259 | -104.2406
11.1296 | -104.2467
11.1333 | -104.2551
11.137 | -104.2699
11.1407 | -104.2898
11.1444 | -104.3063
11.1481 | -104.3208
11.1519 | -104.3399
11.1556 | -104.3585
11.1593 | -104.3785
11.163 | -104.398
11.1667 | -104.4148
11.1704 | -104.4275
11.1741 | -104.4367
11.1778 | -104.4421
11.1815 | -104.4442
11.1852 | -104.4404
11.1889 | -104.434
11.1926 | -104.4255
11.1963 | -104.406
11.2 | -104.3858
11.2037 | -104.3649
11.2074 | -104.3372
11.2111 | -104.3119
11.2148 | -104.2867
11.2185 | -104.2647
11.2222 | -104.2427
11.2259 | -104.2263
11.2296 | -104.2082
11.2333 | -104.1894
11.237 | -104.1762
11.2407 | -104.1613
11.2444 | -104.1467
11.2481 | -104.1339
11.2519 | -104.1245
11.2556 | -104.1218
11.2593 | -104.1223
11.263 | -104.1251
11.2667 | -104.1285
11.2704 | -104.1281
11.2741 | -104.1259
11.2778 | -104.1257
11.2815 | -104.1259
11.2852 | -104.1307
11.2889 | -104.1361
11.2926 | -104.142
11.2963 | -104.1442
11.3 | -104.1419
11.3037 | -104.1393
11.3074 | -104.1401
11.3111 | -104.1398
11.3148 | -104.1343
11.3185 | -104.1306
11.3222 | -104.1267
11.3259 | -104.1221
11.3296 | -104.1188
11.3333 | -104.1181
11.337 | -104.117
11.3407 | -104.1141
11.3444 | -104.1098
11.3481 | -104.1071
11.3519 | -104.1016
11.3556 | -104.0967
11.3593 | -104.0978
11.363 | -104.0943
11.3667 | -104.093
11.3704 | -104.0926
11.3741 | -104.0883
11.3778 | -104.0879
11.3815 | -104.0931
11.3852 | -104.0981
11.3889 | -104.1018
11.3926 | -104.1063
11.3963 | -104.1082
11.4 | -104.1134
11.4037 | -104.1144
11.4074 | -104.1172
11.4111 | -104.1191
11.4148 | -104.1178
11.4185 | -104.1123
11.4222 | -104.1028
11.4259 | -104.0911
11.4296 | -104.0819
11.4333 | -104.0685
11.437 | -104.0564
11.4407 | -104.0441
11.4444 | -104.0261
11.4481 | -104.0074
11.4519 | -103.9872
11.4556 | -103.9655
11.4593 | -103.9438
11.463 | -103.9255
11.4667 | -103.9131
11.4704 | -103.9027
11.4741 | -103.8983
11.4778 | -103.9
11.4815 | -103.9042
11.4852 | -103.9082
11.4889 | -103.9167
11.4926 | -103.9303
11.4963 | -103.9445
11.5 | -103.96
11.5037 | -103.9787
11.5074 | -104.0007
11.5111 | -104.0257
11.5148 | -104.0503
11.5185 | -104.0733
11.5222 | -104.0954
11.5259 | -104.1156
11.5296 | -104.1357
11.5333 | -104.1553
11.537 | -104.1779
11.5407 | -104.2013
11.5444 | -104.2216
11.5481 | -104.2381
11.5519 | -104.2534
11.5556 | -104.2745
11.5593 | -104.2901
11.563 | -104.3083
11.5667 | -104.3231
11.5704 | -104.3375
11.5741 | -104.3478
11.5778 | -104.3574
11.5815 | -104.3677
11.5852 | -104.3759
11.5889 | -104.3787
11.5926 | -104.3778
11.5963 | -104.3774
11.6 | -104.3748
11.6037 | -104.3697
11.6074 | -104.3633
11.6111 | -104.3574
11.6148 | -104.3473
11.6185 | -104.3372
11.6222 | -104.3281
11.6259 | -104.3205
11.6296 | -104.3137
11.6333 | -104.3031
11.637 | -104.2921
11.6407 | -104.2788
11.6444 | -104.2708
11.6481 | -104.2663
11.6519 | -104.2625
11.6556 | -104.2591
11.6593 | -104.2555
11.663 | -104.2531
11.6667 | -104.2519
11.6704 | -104.2491
11.6741 | -104.2415
11.6778 | -104.2335
11.6815 | -104.226
11.6852 | -104.2155
11.6889 | -104.2019
11.6926 | -104.1869
11.6963 | -104.1684
11.7 | -104.1485
11.7037 | -104.1261
11.7074 | -104.1034
11.7111 | -104.0811
11.7148 | -104.0572
11.7185 | -104.0335
11.7222 | -104.0101
11.7259 | -103.9918
11.7296 | -103.9756
11.7333 | -103.9629
11.737 | -103.9529
11.7407 | -103.9465
11.7444 | -103.9453
11.7481 | -103.9479
11.7519 | -103.9708
11.7556 | -103.9823
11.7593 | -103.9956
11.763 | -104.0116
11.7667 | -104.0306
11.7704 | -104.0518
11.7741 | -104.0777
11.7778 | -104.1051
11.7815 | -104.1331
11.7852 | -104.1614
11.7889 | -104.1899
11.7926 | -104.2176
11.7963 | -104.2431
11.8 | -104.2689
11.8037 | -104.2969
11.8074 | -104.3242
11.8111 | -104.3491
11.8148 | -104.3676
11.8185 | -104.3845
11.8222 | -104.4015
11.8259 | -104.4154
11.8296 | -104.4264
11.8333 | -104.438
11.837 | -104.4468
11.8407 | -104.4522
11.8444 | -104.4533
11.8481 | -104.4506
11.8519 | -104.4456
11.8556 | -104.4386
11.8593 | -104.4248
11.863 | -104.4087
11.8667 | -104.3893
11.8704 | -104.3661
11.8741 | -104.3414
11.8778 | -104.3156
11.8815 | -104.2902
11.8852 | -104.2659
11.8889 | -104.2391
11.8926 | -104.2081
11.8963 | -104.175
11.9 | -104.145
11.9037 | -104.1136
11.9074 | -104.0845
11.9111 | -104.0552
11.9148 | -104.0156
11.9185 | -103.9848
11.9222 | -103.9572
11.9259 | -103.9333
11.9296 | -103.9101
11.9333 | -103.8897
11.937 | -103.8711
11.9407 | -103.8491
11.9444 | -103.8297
11.9481 | -103.8132
11.9519 | -103.8003
11.9556 | -103.7865
11.9593 | -103.7773
11.963 | -103.773
11.9667 | -103.7703
11.9704 | -103.7685
11.9741 | -103.769
11.9778 | -103.7701
11.9815 | -103.7759
11.9852 | -103.7864
11.9889 | -103.7984
11.9926 | -103.8121
11.9963 | -103.8289
12 | -103.8483
12.0037 | -103.8701
12.0074 | -103.8939
12.0111 | -103.9202
12.0148 | -103.9501
12.0185 | -103.9836
12.0222 | -104.0151
12.0259 | -104.045
12.0296 | -104.0754
12.0333 | -104.1031
12.037 | -104.1266
12.0407 | -104.1491
12.0444 | -104.1658
12.0481 | -104.181
12.0519 | -104.1924
12.0556 | -104.2001
12.0593 | -104.2029
12.063 | -104.2029
12.0667 | -104.2011
12.0704 | -104.1983
12.0741 | -104.1893
12.0778 | -104.1773
12.0815 | -104.1641
12.0852 | -104.155
12.0889 | -104.1507
12.0926 | -104.1356
12.0963 | -104.122
12.1 | -104.1105
12.1037 | -104.0975
12.1074 | -104.0883
12.1111 | -104.0826
12.1148 | -104.0803
12.1185 | -104.082
12.1222 | -104.0846
12.1259 | -104.0889
12.1296 | -104.0985
12.1333 | -104.1089
12.137 | -104.1154
12.1407 | -104.1255
12.1444 | -104.1393
12.1481 | -104.1529
12.1519 | -104.1711
12.1556 | -104.1889
12.1593 | -104.2076
12.163 | -104.2303
12.1667 | -104.2504
12.1704 | -104.2678
12.1741 | -104.2879
12.1778 | -104.3092
12.1815 | -104.3292
12.1852 | -104.3474
12.1889 | -104.3642
12.1926 | -104.381
12.1963 | -104.3976
12.2 | -104.4136
12.2037 | -104.4304
12.2074 | -104.4509
12.2111 | -104.4707
12.2148 | -104.4889
12.2185 | -104.5043
12.2222 | -104.5134
12.2259 | -104.5217
12.2296 | -104.5277
12.2333 | -104.5314
12.237 | -104.5361
12.2407 | -104.5434
12.2444 | -104.5475
12.2481 | -104.5452
12.2519 | -104.5416
12.2556 | -104.5358
12.2593 | -104.5261
12.263 | -104.5196
12.2667 | -104.5119
12.2704 | -104.5021
12.2741 | -104.489
12.2778 | -104.4777
12.2815 | -104.4668
12.2852 | -104.4549
12.2889 | -104.4543
12.2926 | -104.4466
12.2963 | -104.437
12.3 | -104.4304
12.3037 | -104.4253
12.3074 | -104.4236
12.3111 | -104.4253
12.3148 | -104.4325
12.3185 | -104.4382
12.3222 | -104.446
12.3259 | -104.4523
12.3296 | -104.4587
12.3333 | -104.4651
12.337 | -104.4744
12.3407 | -104.4825
12.3444 | -104.493
12.3481 | -104.5006
12.3519 | -104.5071
12.3556 | -104.5137
12.3593 | -104.5189
12.363 | -104.525
12.3667 | -104.53
12.3704 | -104.5343
12.3741 | -104.5378
12.3778 | -104.5385
12.3815 | -104.5367
12.3852 | -104.5319
12.3889 | -104.5267
12.3926 | -104.5193
12.3963 | -104.5157
12.4 | -104.5123
12.4037 | -104.5046
12.4074 | -104.4967
12.4111 | -104.4906
12.4148 | -104.4846
12.4185 | -104.4729
12.4222 | -104.4617
12.4259 | -104.4573
12.4296 | -104.454
12.4333 | -104.4513
12.437 | -104.4459
12.4407 | -104.4385
12.4444 | -104.4339
12.4481 | -104.4302
12.4519 | -104.4277
12.4556 | -104.4286
12.4593 | -104.4273
12.463 | -104.426
12.4667 | -104.4253
12.4704 | -104.423
12.4741 | -104.4249
12.4778 | -104.4269
12.4815 | -104.4275
12.4852 | -104.431
12.4889 | -104.4334
12.4926 | -104.4377
12.4963 | -104.4427
12.5 | -104.4512
12.5037 | -104.4598
12.5074 | -104.4718
12.5111 | -104.484
12.5148 | -104.496
12.5185 | -104.5101
12.5222 | -104.5243
12.5259 | -104.5386
12.5296 | -104.5552
12.5333 | -104.569
12.537 | -104.5804
12.5407 | -104.595
12.5444 | -104.6079
12.5481 | -104.618
12.5519 | -104.6265
12.5556 | -104.6347
12.5593 | -104.6412
12.563 | -104.645
12.5667 | -104.6504
12.5704 | -104.6557
12.5741 | -104.6581
12.5778 | -104.6557
12.5815 | -104.6552
12.5852 | -104.6571
12.5889 | -104.6573
12.5926 | -104.6536
12.5963 | -104.6492
12.6 | -104.6489
12.6037 | -104.6495
12.6074 | -104.6498
12.6111 | -104.6498
12.6148 | -104.6469
12.6185 | -104.647
12.6222 | -104.6485
12.6259 | -104.6461
12.6296 | -104.6411
12.6333 | -104.6364
12.637 | -104.6343
12.6407 | -104.6348
12.6444 | -104.6326
12.6481 | -104.6282
12.6519 | -104.6262
12.6556 | -104.6248
12.6593 | -104.6225
12.663 | -104.6197
12.6667 | -104.6193
12.6704 | -104.6164
12.6741 | -104.6142
12.6778 | -104.6121
12.6815 | -104.6101
12.6852 | -104.6089
12.6889 | -104.6113
12.6926 | -104.6131
12.6963 | -104.6164
12.7 | -104.6188
12.7037 | -104.6241
12.7074 | -104.6292
12.7111 | -104.6367
12.7148 | -104.6467
12.7185 | -104.6574
12.7222 | -104.668
12.7259 | -104.6801
12.7296 | -104.6972
12.7333 | -104.7136
12.737 | -104.7308
12.7407 | -104.7509
12.7444 | -104.7704
12.7481 | -104.7849
12.7519 | -104.7982
12.7556 | -104.8127
12.7593 | -104.827
12.763 | -104.8399
12.7667 | -104.8487
12.7704 | -104.8547
12.7741 | -104.8581
12.7778 | -104.8564
12.7815 | -104.8526
12.7852 | -104.8477
12.7889 | -104.8373
12.7926 | -104.8231
12.7963 | -104.8132
12.8 | -104.8044
12.8037 | -104.7899
12.8074 | -104.7739
12.8111 | -104.7557
12.8148 | -104.7368
12.8185 | -104.7111
12.8222 | -104.6871
12.8259 | -104.6629
12.8296 | -104.6419
12.8333 | -104.62
12.837 | -104.5954
12.8407 | -104.5734
12.8444 | -104.5567
12.8481 | -104.5381
12.8519 | -104.5172
12.8556 | -104.502
12.8593 | -104.4859
12.863 | -104.4705
12.8667 | -104.4595
12.8704 | -104.4522
12.8741 | -104.4462
12.8778 | -104.4396
12.8815 | -104.4347
12.8852 | -104.4311
12.8889 | -104.4277
12.8926 | -104.4283
12.8963 | -104.4294
12.9 | -104.4309
12.9037 | -104.4296
12.9074 | -104.4147
12.9111 | -104.4091
12.9148 | -104.4062
12.9185 | -104.4054
12.9222 | -104.4025
12.9259 | -104.3992
12.9296 | -104.3972
12.9333 | -104.3944
12.937 | -104.3909
12.9407 | -104.3832
12.9444 | -104.3732
12.9481 | -104.3591
12.9519 | -104.3452
12.9556 | -104.3317
12.9593 | -104.3185
12.963 | -104.3046
12.9667 | -104.2886
12.9704 | -104.2711
12.9741 | -104.2549
12.9778 | -104.2425
12.9815 | -104.2281
12.9852 | -104.2157
12.9889 | -104.2104
12.9926 | -104.2092
12.9963 | -104.2115
13 | -104.2162
13.0037 | -104.2259
13.0074 | -104.2378
13.0111 | -104.2528
13.0148 | -104.2692
13.0185 | -104.288
13.0222 | -104.3087
13.0259 | -104.332
13.0296 | -104.3517
13.0333 | -104.3713
13.037 | -104.393
13.0407 | -104.4152
13.0444 | -104.4354
13.0481 | -104.4558
13.0519 | -104.4779
13.0556 | -104.4936
13.0593 | -104.5064
13.063 | -104.5165
13.0667 | -104.5273
13.0704 | -104.5354
13.0741 | -104.5417
13.0778 | -104.545
13.0815 | -104.5439
13.0852 | -104.5398
13.0889 | -104.5339
13.0926 | -104.5299
13.0963 | -104.5317
13.1 | -104.52
13.1037 | -104.5044
13.1074 | -104.4873
13.1111 | -104.4715
13.1148 | -104.4539
13.1185 | -104.4387
13.1222 | -104.4284
13.1259 | -104.42
13.1296 | -104.4156
13.1333 | -104.4127
13.137 | -104.4114
13.1407 | -104.4086
13.1444 | -104.4063
13.1481 | -104.4059
13.1519 | -104.4037
13.1556 | -104.4069
13.1593 | -104.4128
13.163 | -104.4152
13.1667 | -104.4175
13.1704 | -104.4208
13.1741 | -104.4207
13.1778 | -104.4191
13.1815 | -104.4152
13.1852 | -104.4127
13.1889 | -104.4076
13.1926 | -104.402
13.1963 | -104.3948
13.2 | -104.3874
13.2037 | -104.3807
13.2074 | -104.3731
13.2111 | -104.3634
13.2148 | -104.3514
13.2185 | -104.3381
13.2222 | -104.3217
13.2259 | -104.3069
13.2296 | -104.2918
13.2333 | -104.2743
13.237 | -104.2566
13.2407 | -104.238
13.2444 | -104.2239
13.2481 | -104.2111
13.2519 | -104.1978
13.2556 | -104.1862
13.2593 | -104.1771
13.263 | -104.1697
13.2667 | -104.1668
13.2704 | -104.1665
13.2741 | -104.1713
13.2778 | -104.1789
13.2815 | -104.1914
13.2852 | -104.2086
13.2889 | -104.228
13.2926 | -104.2447
13.2963 | -104.2627
13.3 | -104.2845
13.3037 | -104.309
13.3074 | -104.331
13.3111 | -104.3551
13.3148 | -104.3778
13.3185 | -104.4006
13.3222 | -104.4232
13.3259 | -104.4383
13.3296 | -104.4512
13.3333 | -104.4643
13.337 | -104.4757
13.3407 | -104.4866
13.3444 | -104.4888
13.3481 | -104.4879
13.3519 | -104.486
13.3556 | -104.4665
13.3593 | -104.4587
13.363 | -104.4465
13.3667 | -104.4348
13.3704 | -104.4224
13.3741 | -104.4122
13.3778 | -104.4003
13.3815 | -104.3879
13.3852 | -104.3728
13.3889 | -104.3595
13.3926 | -104.3477
13.3963 | -104.3388
13.4 | -104.3291
13.4037 | -104.3214
13.4074 | -104.3154
13.4111 | -104.3108
13.4148 | -104.3081
13.4185 | -104.3063
13.4222 | -104.3049
13.4259 | -104.3061
13.4296 | -104.3042
13.4333 | -104.3043
13.437 | -104.304
13.4407 | -104.3038
13.4444 | -104.3079
13.4481 | -104.3124
13.4519 | -104.319
13.4556 | -104.3259
13.4593 | -104.3298
13.463 | -104.3374
13.4667 | -104.3418
13.4704 | -104.3444
13.4741 | -104.3497
13.4778 | -104.3578
13.4815 | -104.3668
13.4852 | -104.3794
13.4889 | -104.3923
13.4926 | -104.4023
13.4963 | -104.4124
13.5 | -104.4242
13.5037 | -104.4328
13.5074 | -104.4438
13.5111 | -104.4571
13.5148 | -104.4714
13.5185 | -104.4921
13.5222 | -104.5112
13.5259 | -104.5258
13.5296 | -104.5412
13.5333 | -104.559
13.537 | -104.5791
13.5407 | -104.5983
13.5444 | -104.622
13.5481 | -104.6453
13.5519 | -104.6686
13.5556 | -104.6903
13.5593 | -104.7126
13.563 | -104.731
13.5667 | -104.7487
13.5704 | -104.7666
13.5741 | -104.7828
13.5778 | -104.7981
13.5815 | -104.8107
13.5852 | -104.8201
13.5889 | -104.8281
13.5926 | -104.8326
13.5963 | -104.8404
13.6 | -104.8466
13.6037 | -104.8487
13.6074 | -104.8522
13.6111 | -104.8513
13.6148 | -104.8496
13.6185 | -104.8477
13.6222 | -104.8456
13.6259 | -104.8493
13.6296 | -104.8466
13.6333 | -104.8435
13.637 | -104.8422
13.6407 | -104.8382
13.6444 | -104.8332
13.6481 | -104.8274
13.6519 | -104.8232
13.6556 | -104.8236
13.6593 | -104.82
13.663 | -104.8175
13.6667 | -104.8191
13.6704 | -104.8192
13.6741 | -104.8155
13.6778 | -104.8116
13.6815 | -104.8091
13.6852 | -104.8072
13.6889 | -104.8083
13.6926 | -104.8102
13.6963 | -104.8093
13.7 | -104.8085
13.7037 | -104.8078
13.7074 | -104.8074
13.7111 | -104.8045
13.7148 | -104.8017
13.7185 | -104.8015
13.7222 | -104.8021
13.7259 | -104.7998
13.7296 | -104.7941
13.7333 | -104.7897
13.737 | -104.7892
13.7407 | -104.7879
13.7444 | -104.7863
13.7481 | -104.7861
13.7519 | -104.7879
13.7556 | -104.7894
13.7593 | -104.7921
13.763 | -104.7931
13.7667 | -104.7923
13.7704 | -104.7901
13.7741 | -104.7871
13.7778 | -104.7867
13.7815 | -104.7842
13.7852 | -104.7782
13.7889 | -104.7696
13.7926 | -104.7606
13.7963 | -104.7512
13.8 | -104.7436
13.8037 | -104.7319
13.8074 | -104.7211
13.8111 | -104.7095
13.8148 | -104.6984
13.8185 | -104.687
13.8222 | -104.6703
13.8259 | -104.659
13.8296 | -104.6467
13.8333 | -104.632
13.837 | -104.6195
13.8407 | -104.6068
13.8444 | -104.5936
13.8481 | -104.5813
13.8519 | -104.5703
13.8556 | -104.5623
13.8593 | -104.5563
13.863 | -104.5501
13.8667 | -104.5431
13.8704 | -104.5371
13.8741 | -104.5306
13.8778 | -104.5255
13.8815 | -104.5207
13.8852 | -104.5158
13.8889 | -104.5199
13.8926 | -104.5181
13.8963 | -104.5143
13.9 | -104.4944
13.9037 | -104.4812
13.9074 | -104.471
13.9111 | -104.4629
13.9148 | -104.4531
13.9185 | -104.4404
13.9222 | -104.4273
13.9259 | -104.413
13.9296 | -104.3993
13.9333 | -104.3837
13.937 | -104.3654
13.9407 | -104.3446
13.9444 | -104.3239
13.9481 | -104.3025
13.9519 | -104.2795
13.9556 | -104.2571
13.9593 | -104.2331
13.963 | -104.2078
13.9667 | -104.1826
13.9704 | -104.1605
13.9741 | -104.1414
13.9778 | -104.1226
13.9815 | -104.1062
13.9852 | -104.0941
13.9889 | -104.0858
13.9926 | -104.0811
13.9963 | -104.0811
14 | -104.0829
14.0037 | -104.09
14.0074 | -104.1024
14.0111 | -104.1179
14.0148 | -104.1389
14.0185 | -104.1636
14.0222 | -104.1953
14.0259 | -104.2296
14.0296 | -104.2662
14.0333 | -104.3063
14.037 | -104.3464
14.0407 | -104.39
14.0444 | -104.4358
14.0481 | -104.4808
14.0519 | -104.5265
14.0556 | -104.5725
14.0593 | -104.6178
14.063 | -104.6625
14.0667 | -104.7073
14.0704 | -104.7541
14.0741 | -104.7975
14.0778 | -104.8371
14.0815 | -104.8744
14.0852 | -104.906
14.0889 | -104.9381
14.0926 | -104.9687
14.0963 | -105.0003
14.1 | -105.0348
14.1037 | -105.0777
14.1074 | -105.1021
14.1111 | -105.12
14.1148 | -105.1358
14.1185 | -105.1511
14.1222 | -105.1654
14.1259 | -105.175
14.1296 | -105.1823
14.1333 | -105.186
14.137 | -105.188
14.1407 | -105.1865
14.1444 | -105.1843
14.1481 | -105.1825
14.1519 | -105.1794
14.1556 | -105.1752
14.1593 | -105.1694
14.163 | -105.1659
14.1667 | -105.1625
14.1704 | -105.1581
14.1741 | -105.1496
14.1778 | -105.1415
14.1815 | -105.1318
14.1852 | -105.121
14.1889 | -105.1104
14.1926 | -105.1004
14.1963 | -105.0836
14.2 | -105.0714
14.2037 | -105.0637
14.2074 | -105.0542
14.2111 | -105.0447
14.2148 | -105.0346
14.2185 | -105.0222
14.2222 | -105.0108
14.2259 | -104.9976
14.2296 | -104.987
14.2333 | -104.9756
14.237 | -104.9627
14.2407 | -104.9513
14.2444 | -104.9416
14.2481 | -104.9312
14.2519 | -104.9197
14.2556 | -104.9078
14.2593 | -104.8991
14.263 | -104.8915
14.2667 | -104.8814
14.2704 | -104.8703
14.2741 | -104.8606
14.2778 | -104.853
14.2815 | -104.8405
14.2852 | -104.8288
14.2889 | -104.8215
14.2926 | -104.8111
14.2963 | -104.7976
14.3 | -104.7835
14.3037 | -104.7708
14.3074 | -104.7614
14.3111 | -104.7521
14.3148 | -104.7377
14.3185 | -104.7235
14.3222 | -104.7091
14.3259 | -104.6956
14.3296 | -104.6816
14.3333 | -104.6642
14.337 | -104.646
14.3407 | -104.6296
14.3444 | -104.6132
14.3481 | -104.5959
14.3519 | -104.5768
14.3556 | -104.5572
14.3593 | -104.5383
14.363 | -104.5227
14.3667 | -104.507
14.3704 | -104.4915
14.3741 | -104.4787
14.3778 | -104.4684
14.3815 | -104.4582
14.3852 | -104.4507
14.3889 | -104.4454
14.3926 | -104.4427
14.3963 | -104.442
14.4 | -104.4472
14.4037 | -104.4521
14.4074 | -104.4582
14.4111 | -104.4697
14.4148 | -104.4835
14.4185 | -104.5012
14.4222 | -104.5239
14.4259 | -104.5589
14.4296 | -104.5834
14.4333 | -104.61
14.437 | -104.638
14.4407 | -104.6688
14.4444 | -104.7009
14.4481 | -104.7323
14.4519 | -104.7636
14.4556 | -104.7948
14.4593 | -104.8261
14.463 | -104.8576
14.4667 | -104.8877
14.4704 | -104.9176
14.4741 | -104.9477
14.4778 | -104.9757
14.4815 | -104.9993
14.4852 | -105.022
14.4889 | -105.0462
14.4926 | -105.0672
14.4963 | -105.0801
14.5 | -105.0895
14.5037 | -105.0977
14.5074 | -105.1046
14.5111 | -105.1054
14.5148 | -105.1039
14.5185 | -105.1012
14.5222 | -105.0968
14.5259 | -105.0881
14.5296 | -105.0773
14.5333 | -105.0621
14.537 | -105.0477
14.5407 | -105.0364
14.5444 | -105.0245
14.5481 | -105.0089
14.5519 | -104.9946
14.5556 | -104.9806
14.5593 | -104.9681
14.563 | -104.9539
14.5667 | -104.9423
14.5704 | -104.9294
14.5741 | -104.9196
14.5778 | -104.9112
14.5815 | -104.9032
14.5852 | -104.8973
14.5889 | -104.8936
14.5926 | -104.8915
14.5963 | -104.8899
14.6 | -104.8896
14.6037 | -104.8891
14.6074 | -104.8887
14.6111 | -104.8899
14.6148 | -104.8913
14.6185 | -104.8927
14.6222 | -104.894
14.6259 | -104.8952
14.6296 | -104.8944
14.6333 | -104.8945
14.637 | -104.8945
14.6407 | -104.8924
14.6444 | -104.8926
14.6481 | -104.8913
14.6519 | -104.8913
14.6556 | -104.8903
14.6593 | -104.8885
14.663 | -104.8847
14.6667 | -104.8818
14.6704 | -104.8817
14.6741 | -104.8807
14.6778 | -104.8782
14.6815 | -104.8749
14.6852 | -104.8726
14.6889 | -104.8732
14.6926 | -104.875
14.6963 | -104.8757
14.7 | -104.876
14.7037 | -104.8753
14.7074 | -104.8789
14.7111 | -104.8798
14.7148 | -104.8828
14.7185 | -104.8857
14.7222 | -104.8887
14.7259 | -104.8928
14.7296 | -104.8986
14.7333 | -104.9034
14.737 | -104.9053
14.7407 | -104.9064
14.7444 | -104.9106
14.7481 | -104.913
14.7519 | -104.9142
14.7556 | -104.9126
14.7593 | -104.9087
14.763 | -104.9062
14.7667 | -104.905
14.7704 | -104.9022
14.7741 | -104.8994
14.7778 | -104.8956
14.7815 | -104.8927
14.7852 | -104.8886
14.7889 | -104.8864
14.7926 | -104.8817
14.7963 | -104.8756
14.8 | -104.8675
14.8037 | -104.8616
14.8074 | -104.8546
14.8111 | -104.8481
14.8148 | -104.8452
14.8185 | -104.8436
14.8222 | -104.8406
14.8259 | -104.8381
14.8296 | -104.8362
14.8333 | -104.8357
14.837 | -104.8356
14.8407 | -104.8399
14.8444 | -104.843
14.8481 | -104.8465
14.8519 | -104.8497
14.8556 | -104.8505
14.8593 | -104.8521
14.863 | -104.853
14.8667 | -104.8582
14.8704 | -104.8655
14.8741 | -104.8702
14.8778 | -104.8751
14.8815 | -104.8804
14.8852 | -104.8883
14.8889 | -104.8953
14.8926 | -104.8948
14.8963 | -104.8816
14.9 | -104.8763
14.9037 | -104.8713
14.9074 | -104.8662
14.9111 | -104.8607
14.9148 | -104.8532
14.9185 | -104.842
14.9222 | -104.8326
14.9259 | -104.8186
14.9296 | -104.8015
14.9333 | -104.7845
14.937 | -104.7645
14.9407 | -104.7409
14.9444 | -104.7174
14.9481 | -104.6975
14.9519 | -104.6774
14.9556 | -104.6539
14.9593 | -104.6314
14.963 | -104.6105
14.9667 | -104.588
14.9704 | -104.5684
14.9741 | -104.5511
14.9778 | -104.5332
14.9815 | -104.5153
14.9852 | -104.5014
14.9889 | -104.4911
14.9926 | -104.4819
14.9963 | -104.4734
15 | -104.4659
15.0037 | -104.4619
15.0074 | -104.4636
15.0111 | -104.4665
15.0148 | -104.4727
15.0185 | -104.4799
15.0222 | -104.4886
15.0259 | -104.4958
15.0296 | -104.5036
15.0333 | -104.5153
15.037 | -104.5301
15.0407 | -104.5463
15.0444 | -104.5646
15.0481 | -104.5839
15.0519 | -104.6048
15.0556 | -104.6243
15.0593 | -104.6446
15.063 | -104.6653
15.0667 | -104.6898
15.0704 | -104.7134
15.0741 | -104.7346
15.0778 | -104.7555
15.0815 | -104.775
15.0852 | -104.7926
15.0889 | -104.8076
15.0926 | -104.8248
15.0963 | -104.8398
15.1 | -104.8531
15.1037 | -104.8678
15.1074 | -104.8896
15.1111 | -104.9182
15.1148 | -104.9316
15.1185 | -104.9415
15.1222 | -104.95
15.1259 | -104.9584
15.1296 | -104.9662
15.1333 | -104.9715
15.137 | -104.9725
15.1407 | -104.9755
15.1444 | -104.9786
15.1481 | -104.9821
15.1519 | -104.9859
15.1556 | -104.9894
15.1593 | -104.9887
15.163 | -104.985
15.1667 | -104.9794
15.1704 | -104.9752
15.1741 | -104.9704
15.1778 | -104.9649
15.1815 | -104.9587
15.1852 | -104.9546
15.1889 | -104.9472
15.1926 | -104.9392
15.1963 | -104.9329
15.2 | -104.9277
15.2037 | -104.9232
15.2074 | -104.9177
15.2111 | -104.9123
15.2148 | -104.9108
15.2185 | -104.9121
15.2222 | -104.9096
15.2259 | -104.9039
15.2296 | -104.9033
15.2333 | -104.9051
15.237 | -104.9066
15.2407 | -104.9102
15.2444 | -104.9174
15.2481 | -104.9261
15.2519 | -104.9311
15.2556 | -104.9363
15.2593 | -104.9444
15.263 | -104.9533
15.2667 | -104.9617
15.2704 | -104.9715
15.2741 | -104.982
15.2778 | -104.9968
15.2815 | -105.0104
15.2852 | -105.0182
15.2889 | -105.0265
15.2926 | -105.0383
15.2963 | -105.0492
15.3 | -105.0598
15.3037 | -105.0725
15.3074 | -105.087
15.3111 | -105.0974
15.3148 | -105.1075
15.3185 | -105.1166
15.3222 | -105.1251
15.3259 | -105.1306
15.3296 | -105.1353
15.3333 | -105.1387
15.337 | -105.1445
15.3407 | -105.1511
15.3444 | -105.153
15.3481 | -105.1541
15.3519 | -105.1577
15.3556 | -105.1621
15.3593 | -105.1598
15.363 | -105.1566
15.3667 | -105.1502
15.3704 | -105.1431
15.3741 | -105.1362
15.3778 | -105.1282
15.3815 | -105.1168
15.3852 | -105.1066
15.3889 | -105.0942
15.3926 | -105.0801
15.3963 | -105.0652
15.4 | -105.048
15.4037 | -105.0285
15.4074 | -105.0126
15.4111 | -104.9952
15.4148 | -104.9721
15.4185 | -104.95
15.4222 | -104.9265
15.4259 | -104.9065
15.4296 | -104.8878
15.4333 | -104.8664
15.437 | -104.8463
15.4407 | -104.8309
15.4444 | -104.8172
15.4481 | -104.8042
15.4519 | -104.793
15.4556 | -104.7839
15.4593 | -104.7779
15.463 | -104.7728
15.4667 | -104.7678
15.4704 | -104.7684
15.4741 | -104.7711
15.4778 | -104.7746
15.4815 | -104.7802
15.4852 | -104.788
15.4889 | -104.7965
15.4926 | -104.8157
15.4963 | -104.8254
15.5 | -104.8352
15.5037 | -104.8477
15.5074 | -104.8596
15.5111 | -104.296
15.5148 | -103.2574
15.5185 | -102.9217
15.5222 | -102.6342
15.5259 | -102.3276
15.5296 | -102.0042
15.5333 | -101.6656
15.537 | -101.3147
15.5407 | -100.9545
15.5444 | -100.5899
15.5481 | -100.2245
15.5519 | -99.8601
15.5556 | -99.5002
15.5593 | -99.1459
15.563 | -98.801
15.5667 | -98.4688
15.5704 | -98.1498
15.5741 | -97.8451
15.5778 | -97.557
15.5815 | -97.2876
15.5852 | -97.0368
15.5889 | -96.8058
15.5926 | -96.5946
15.5963 | -96.4046
15.6 | -96.2368
15.6037 | -96.0912
15.6074 | -95.968
15.6111 | -95.8679
15.6148 | -95.7911
15.6185 | -95.7375
15.6222 | -95.7075
15.6259 | -95.701
15.6296 | -95.7175
15.6333 | -95.7574
15.637 | -95.8213
15.6407 | -95.9082
15.6444 | -96.0178
15.6481 | -96.1497
15.6519 | -96.3039
15.6556 | -96.4804
15.6593 | -96.6781
15.663 | -96.8963
15.6667 | -97.1341
15.6704 | -97.391
15.6741 | -97.6654
15.6778 | -97.9565
15.6815 | -98.2633
15.6852 | -98.5843
15.6889 | -98.9175
15.6926 | -99.2598
15.6963 | -99.6098
15.7 | -99.9647
15.7037 | -100.3223
15.7074 | -100.6803
15.7111 | -101.0342
15.7148 | -101.3806
15.7185 | -101.7173
15.7222 | -102.0431
15.7259 | -102.3534
15.7296 | -102.646
15.7333 | -102.9188
15.737 | -103.5268
15.7407 | -104.5624
15.7444 | -104.687
15.7481 | -104.6924
15.7519 | -104.6961
15.7556 | -104.6997
15.7593 | -104.7037
15.763 | -104.7082
15.7667 | -104.715
15.7704 | -104.7206
15.7741 | -104.7272
15.7778 | -104.7328
15.7815 | -104.738
15.7852 | -104.7473
15.7889 | -104.7548
15.7926 | -104.7608
15.7963 | -104.7685
15.8 | -104.7769
15.8037 | -104.7845
15.8074 | -104.7887
15.8111 | -104.7913
15.8148 | -104.7963
15.8185 | -104.8028
15.8222 | -104.8097
15.8259 | -104.8153
15.8296 | -104.8197
15.8333 | -104.8245
15.837 | -104.8271
15.8407 | -104.8295
15.8444 | -104.8318
15.8481 | -104.833
15.8519 | -104.8369
15.8556 | -104.8411
15.8593 | -104.8438
15.863 | -104.8445
15.8667 | -104.8479
15.8704 | -104.853
15.8741 | -104.8568
15.8778 | -104.8573
15.8815 | -104.8544
15.8852 | -104.8517
15.8889 | -104.844
15.8926 | -104.8385
15.8963 | -104.8317
15.9 | -104.8249
15.9037 | -104.8192
15.9074 | -104.8095
15.9111 | -104.7975
15.9148 | -104.7857
15.9185 | -104.7768
15.9222 | -104.7656
15.9259 | -104.7553
15.9296 | -104.7447
15.9333 | -104.7324
15.937 | -104.7206
15.9407 | -104.7073
15.9444 | -104.6952
15.9481 | -104.6815
15.9519 | -104.6681
15.9556 | -104.6562
15.9593 | -104.6471
15.963 | -104.6404
15.9667 | -104.6322
15.9704 | -104.6258
15.9741 | -104.621
15.9778 | -104.6162
15.9815 | -104.6116
15.9852 | -104.6095
15.9889 | -104.6112
15.9926 | -104.6111
15.9963 | -104.6112
16 | -104.6071
16.0037 | -104.6051
16.0074 | -104.6059
16.0111 | -104.6079
16.0148 | -104.6113
16.0185 | -104.6157
16.0222 | -104.6199
16.0259 | -104.6235
16.0296 | -104.623
16.0333 | -104.6238
16.037 | -104.6255
16.0407 | -104.6283
16.0444 | -104.6283
16.0481 | -104.628
16.0519 | -104.6261
16.0556 | -104.6235
16.0593 | -104.6196
16.063 | -104.6135
16.0667 | -104.6073
16.0704 | -104.6007
16.0741 | -104.5949
16.0778 | -104.5869
16.0815 | -104.575
16.0852 | -104.5645
16.0889 | -104.5547
16.0926 | -104.543
16.0963 | -104.5281
16.1 | -104.516
16.1037 | -104.509
16.1074 | -104.5038
16.1111 | -104.4989
16.1148 | -104.4921
16.1185 | -104.4909
16.1222 | -104.4895
16.1259 | -104.4881
16.1296 | -104.4874
16.1333 | -104.4858
16.137 | -104.4857
16.1407 | -104.4869
16.1444 | -104.4858
16.1481 | -104.4862
16.1519 | -104.4907
16.1556 | -104.4961
16.1593 | -104.5017
16.163 | -104.5089
16.1667 | -104.5176
16.1704 | -104.5277
16.1741 | -104.5342
16.1778 | -104.5414
16.1815 | -104.5509
16.1852 | -104.5583
16.1889 | -104.5668
16.1926 | -104.5754
16.1963 | -104.5823
16.2 | -104.5887
16.2037 | -104.5952
16.2074 | -104.6023
16.2111 | -104.611
16.2148 | -104.6189
16.2185 | -104.6244
16.2222 | -104.6334
16.2259 | -104.6419
16.2296 | -104.6507
16.2333 | -104.6581
16.237 | -104.6643
16.2407 | -104.6663
16.2444 | -104.6674
16.2481 | -104.6676
16.2519 | -104.6671
16.2556 | -104.668
16.2593 | -104.6665
16.263 | -104.6661
16.2667 | -104.6638
16.2704 | -104.6599
16.2741 | -104.6526
16.2778 | -104.6445
16.2815 | -104.6368
16.2852 | -104.6288
16.2889 | -104.6219
16.2926 | -104.6152
16.2963 | -104.6072
16.3 | -104.6002
16.3037 | -104.5917
16.3074 | -104.5808
16.3111 | -104.5692
16.3148 | -104.5617
16.3185 | -104.5559
16.3222 | -104.5498
16.3259 | -104.5418
16.3296 | -104.5368
16.3333 | -104.5293
16.337 | -104.518
16.3407 | -104.5072
16.3444 | -104.5001
16.3481 | -104.4964
16.3519 | -104.4926
16.3556 | -104.4874
16.3593 | -104.4837
16.363 | -104.4806
16.3667 | -104.479
16.3704 | -104.48
16.3741 | -104.4814
16.3778 | -104.4862
16.3815 | -104.4906
16.3852 | -104.4927
16.3889 | -104.4951
16.3926 | -104.4987
16.3963 | -104.5043
16.4 | -104.5123
16.4037 | -104.5216
16.4074 | -104.5338
16.4111 | -104.5465
16.4148 | -104.5577
16.4185 | -104.5721
16.4222 | -104.5871
16.4259 | -104.602
16.4296 | -104.62
16.4333 | -104.6406
16.437 | -104.6612
16.4407 | -104.6807
16.4444 | -104.6989
16.4481 | -104.719
16.4519 | -104.7406
16.4556 | -104.7624
16.4593 | -104.7864
16.463 | -104.8109
16.4667 | -104.8336
16.4704 | -104.8545
16.4741 | -104.8755
16.4778 | -104.8966
16.4815 | -104.9149
16.4852 | -104.9321
16.4889 | -104.9477
16.4926 | -104.9646
16.4963 | -104.9804
16.5 | -104.9974
16.5037 | -105.0109
16.5074 | -105.0242
16.5111 | -105.0357
16.5148 | -105.0453
16.5185 | -105.0542
16.5222 | -105.0599
16.5259 | -105.0644
16.5296 | -105.0669
16.5333 | -105.0694
16.537 | -105.0732
16.5407 | -105.0776
16.5444 | -105.0842
16.5481 | -105.0871
16.5519 | -105.0838
16.5556 | -105.083
16.5593 | -105.0849
16.563 | -105.0837
16.5667 | -105.0814
16.5704 | -105.0793
16.5741 | -105.0802
16.5778 | -105.0818
16.5815 | -105.0816
16.5852 | -105.0816
16.5889 | -105.0817
16.5926 | -105.0834
16.5963 | -105.0843
16.6 | -105.0863
16.6037 | -105.0891
16.6074 | -105.0897
16.6111 | -105.0918
16.6148 | -105.0944
16.6185 | -105.0935
16.6222 | -105.0938
16.6259 | -105.0964
16.6296 | -105.0995
16.6333 | -105.0991
16.637 | -105.0962
16.6407 | -105.0955
16.6444 | -105.0959
16.6481 | -105.0958
16.6519 | -105.0952
16.6556 | -105.096
16.6593 | -105.0922
16.663 | -105.087
16.6667 | -105.0823
16.6704 | -105.0767
16.6741 | -105.0721
16.6778 | -105.0675
16.6815 | -105.0631
16.6852 | -105.0548
16.6889 | -105.0477
16.6926 | -105.0382
16.6963 | -105.027
16.7 | -105.0181
16.7037 | -105.0084
16.7074 | -104.9969
16.7111 | -104.9895
16.7148 | -104.9823
16.7185 | -104.9774
16.7222 | -104.9738
16.7259 | -104.9679
16.7296 | -104.9606
16.7333 | -104.9533
16.737 | -104.9456
16.7407 | -104.9409
16.7444 | -104.9377
16.7481 | -104.936
16.7519 | -104.9342
16.7556 | -104.9342
16.7593 | -104.9349
16.763 | -104.9377
16.7667 | -104.9407
16.7704 | -104.9442
16.7741 | -104.9505
16.7778 | -104.9556
16.7815 | -104.9609
16.7852 | -104.9676
16.7889 | -104.9801
16.7926 | -104.9926
16.7963 | -105.0047
16.8 | -105.0187
16.8037 | -105.0342
16.8074 | -105.048
16.8111 | -105.0649
16.8148 | -105.0786
16.8185 | -105.0916
16.8222 | -105.104
16.8259 | -105.1178
16.8296 | -105.1313
16.8333 | -105.144
16.837 | -105.1567
16.8407 | -105.1655
16.8444 | -105.1755
16.8481 | -105.1857
16.8519 | -105.1951
16.8556 | -105.2032
16.8593 | -105.2132
16.863 | -105.2201
16.8667 | -105.2205
16.8704 | -105.2202
16.8741 | -105.2225
16.8778 | -105.223
16.8815 | -105.2109
16.8852 | -105.2
16.8889 | -105.1907
16.8926 | -105.183
16.8963 | -105.1716
16.9 | -105.1635
16.9037 | -105.1553
16.9074 | -105.1433
16.9111 | -105.1298
16.9148 | -105.1128
16.9185 | -105.0982
16.9222 | -105.0847
16.9259 | -105.0719
16.9296 | -105.0568
16.9333 | -105.0428
16.937 | -105.0279
16.9407 | -105.0145
16.9444 | -105.0042
16.9481 | -104.9938
16.9519 | -104.9803
16.9556 | -104.9676
16.9593 | -104.9581
16.963 | -104.9489
16.9667 | -104.9427
16.9704 | -104.9383
16.9741 | -104.9326
16.9778 | -104.9281
16.9815 | -104.925
16.9852 | -104.9241
16.9889 | -104.9226
16.9926 | -104.921
16.9963 | -104.921
17 | -104.9226
17.0037 | -104.9251
17.0074 | -104.9303
17.0111 | -104.936
17.0148 | -104.939
17.0185 | -104.9449
17.0222 | -104.951
17.0259 | -104.9559
17.0296 | -104.9595
17.0333 | -104.9637
17.037 | -104.9683
17.0407 | -104.9713
17.0444 | -104.9746
17.0481 | -104.9793
17.0519 | -104.9822
17.0556 | -104.9882
17.0593 | -104.9936
17.063 | -104.9981
17.0667 | -104.9982
17.0704 | -104.9978
17.0741 | -105
17.0778 | -105.0037
17.0815 | -105.0112
17.0852 | -105.0145
17.0889 | -105.016
17.0926 | -105.0171
17.0963 | -105.019
17.1 | -105.0186
17.1037 | -105.0195
17.1074 | -105.0227
17.1111 | -105.0258
17.1148 | -105.0276
17.1185 | -105.0283
17.1222 | -105.0293
17.1259 | -105.0363
17.1296 | -105.0368
17.1333 | -105.0326
17.137 | -105.0265
17.1407 | -105.0224
17.1444 | -105.0169
17.1481 | -105.0113
17.1519 | -105.0064
17.1556 | -105.0024
17.1593 | -104.9995
17.163 | -104.9955
17.1667 | -104.9881
17.1704 | -104.9822
17.1741 | -104.9747
17.1778 | -104.9672
17.1815 | -104.9626
17.1852 | -104.9548
17.1889 | -104.9459
17.1926 | -104.9391
17.1963 | -104.9345
17.2 | -104.9305
17.2037 | -104.9266
17.2074 | -104.9215
17.2111 | -104.9166
17.2148 | -104.9142
17.2185 | -104.914
17.2222 | -104.9144
17.2259 | -104.914
17.2296 | -104.9126
17.2333 | -104.913
17.237 | -104.9144
17.2407 | -104.9143
17.2444 | -104.9139
17.2481 | -104.9154
17.2519 | -104.917
17.2556 | -104.9173
17.2593 | -104.9179
17.263 | -104.9189
17.2667 | -104.9191
17.2704 | -104.9235
17.2741 | -104.9271
17.2778 | -104.9305
17.2815 | -104.935
17.2852 | -104.9397
17.2889 | -104.9426
17.2926 | -104.9438
17.2963 | -104.9473
17.3 | -104.9484
17.3037 | -104.9489
17.3074 | -104.952
17.3111 | -104.9557
17.3148 | -104.958
17.3185 | -104.9574
17.3222 | -104.9542
17.3259 | -104.9518
17.3296 | -104.9492
17.3333 | -104.9482
17.337 | -104.9466
17.3407 | -104.9418
17.3444 | -104.9364
17.3481 | -104.9329
17.3519 | -104.9268
17.3556 | -104.9186
17.3593 | -104.9113
17.363 | -104.9047
17.3667 | -104.8974
17.3704 | -104.89
17.3741 | -104.8816
17.3778 | -104.8733
17.3815 | -104.8666
17.3852 | -104.8595
17.3889 | -104.8525
17.3926 | -104.8467
17.3963 | -104.8421
17.4 | -104.8384
17.4037 | -104.8351
17.4074 | -104.832
17.4111 | -104.8289
17.4148 | -104.8276
17.4185 | -104.8274
17.4222 | -104.8318
17.4259 | -104.8355
17.4296 | -104.8363
17.4333 | -104.8408
17.437 | -104.846
17.4407 | -104.8517
17.4444 | -104.8577
17.4481 | -104.8649
17.4519 | -104.8695
17.4556 | -104.8757
17.4593 | -104.884
17.463 | -104.891
17.4667 | -104.8995
17.4704 | -104.908
17.4741 | -104.9172
17.4778 | -104.9277
17.4815 | -104.9397
17.4852 | -104.9502
17.4889 | -104.9616
17.4926 | -104.9721
17.4963 | -104.9817
17.5 | -104.991
17.5037 | -105.0016
17.5074 | -105.0121
17.5111 | -105.0198
17.5148 | -105.0302
17.5185 | -105.0373
17.5222 | -105.0441
17.5259 | -105.0495
17.5296 | -105.0566
17.5333 | -105.062
17.537 | -105.0692
17.5407 | -105.0769
17.5444 | -105.078
17.5481 | -105.0807
17.5519 | -105.0864
17.5556 | -105.0889
17.5593 | -105.0902
17.563 | -105.0928
17.5667 | -105.0949
17.5704 | -105.0961
17.5741 | -105.0942
17.5778 | -105.0956
17.5815 | -105.0966
17.5852 | -105.0969
17.5889 | -105.1004
17.5926 | -105.1026
17.5963 | -105.1023
17.6 | -105.1021
17.6037 | -105.1036
17.6074 | -105.1043
17.6111 | -105.1025
17.6148 | -105.1016
17.6185 | -105.0985
17.6222 | -105.0937
17.6259 | -105.089
17.6296 | -105.0807
17.6333 | -105.0755
17.637 | -105.0698
17.6407 | -105.0638
17.6444 | -105.0571
17.6481 | -105.0496
17.6519 | -105.0392
17.6556 | -105.0303
17.6593 | -105.0216
17.663 | -105.0142
17.6667 | -105.008
17.6704 | -104.9994
17.6741 | -104.9885
17.6778 | -104.9781
17.6815 | -104.9728
17.6852 | -104.9664
17.6889 | -104.9559
17.6926 | -104.9452
17.6963 | -104.936
17.7 | -104.9264
17.7037 | -104.919
17.7074 | -104.9134
17.7111 | -104.9059
17.7148 | -104.8986
17.7185 | -104.8908
17.7222 | -104.8845
17.7259 | -104.8773
17.7296 | -104.8685
17.7333 | -104.8594
17.737 | -104.8525
17.7407 | -104.8464
17.7444 | -104.8397
17.7481 | -104.8318
17.7519 | -104.8229
17.7556 | -104.8112
17.7593 | -104.8036
17.763 | -104.7983
17.7667 | -104.79
17.7704 | -104.7794
17.7741 | -104.7697
17.7778 | -104.7606
17.7815 | -104.7514
17.7852 | -104.7435
17.7889 | -104.7342
17.7926 | -104.7257
17.7963 | -104.7193
17.8 | -104.7114
17.8037 | -104.7016
17.8074 | -104.6936
17.8111 | -104.6882
17.8148 | -104.682
17.8185 | -104.6717
17.8222 | -104.6645
17.8259 | -104.6616
17.8296 | -104.6548
17.8333 | -104.649
17.837 | -104.644
17.8407 | -104.6385
17.8444 | -104.6355
17.8481 | -104.6331
17.8519 | -104.6314
17.8556 | -104.63
17.8593 | -104.6286
17.863 | -104.6289
17.8667 | -104.6312
17.8704 | -104.633
17.8741 | -104.6316
17.8778 | -104.6327
17.8815 | -104.6353
17.8852 | -104.6372
17.8889 | -104.639
17.8926 | -104.6412
17.8963 | -104.6416
17.9 | -104.645
17.9037 | -104.648
17.9074 | -104.6505
17.9111 | -104.6495
17.9148 | -104.65
17.9185 | -104.6486
17.9222 | -104.6482
17.9259 | -104.6486
17.9296 | -104.6492
17.9333 | -104.6495
17.937 | -104.646
17.9407 | -104.6404
17.9444 | -104.6399
17.9481 | -104.6395
17.9519 | -104.6384
17.9556 | -104.6366
17.9593 | -104.6322
17.963 | -104.6265
17.9667 | -104.6228
17.9704 | -104.6196
17.9741 | -104.6185
17.9778 | -104.6172
17.9815 | -104.6149
17.9852 | -104.6118
17.9889 | -104.61
17.9926 | -104.6077
17.9963 | -104.605
18 | -104.6033
18.0037 | -104.6023
18.0074 | -104.6014
18.0111 | -104.6022
18.0148 | -104.5999
18.0185 | -104.5974
18.0222 | -104.5976
18.0259 | -104.5997
18.0296 | -104.5984
18.0333 | -104.599
18.037 | -104.6011
18.0407 | -104.6029
18.0444 | -104.6059
18.0481 | -104.6107
18.0519 | -104.6142
18.0556 | -104.6188
18.0593 | -104.6246
18.063 | -104.6281
18.0667 | -104.6324
18.0704 | -104.6375
18.0741 | -104.6486
18.0778 | -104.6572
18.0815 | -104.6643
18.0852 | -104.6748
18.0889 | -104.6819
18.0926 | -104.6864
18.0963 | -104.6904
18.1 | -104.7
18.1037 | -104.7085
18.1074 | -104.7155
18.1111 | -104.7199
18.1148 | -104.7282
18.1185 | -104.7344
18.1222 | -104.7373
18.1259 | -104.7427
18.1296 | -104.7503
18.1333 | -104.7515
18.137 | -104.7505
18.1407 | -104.7494
18.1444 | -104.7485
18.1481 | -104.7439
18.1519 | -104.737
18.1556 | -104.7275
18.1593 | -104.7178
18.163 | -104.711
18.1667 | -104.6934
18.1704 | -104.6854
18.1741 | -104.6745
18.1778 | -104.6635
18.1815 | -104.652
18.1852 | -104.6386
18.1889 | -104.6231
18.1926 | -104.6075
18.1963 | -104.5919
18.2 | -104.5759
18.2037 | -104.5622
18.2074 | -104.5471
18.2111 | -104.5298
18.2148 | -104.5142
18.2185 | -104.5026
18.2222 | -104.4894
18.2259 | -104.4748
18.2296 | -104.4607
18.2333 | -104.4443
18.237 | -104.4281
18.2407 | -104.416
18.2444 | -104.4084
18.2481 | -104.3992
18.2519 | -104.3924
18.2556 | -104.3872
18.2593 | -104.3837
18.263 | -104.3811
18.2667 | -104.3794
18.2704 | -104.3776
18.2741 | -104.3775
18.2778 | -104.3801
18.2815 | -104.3821
18.2852 | -104.3851
18.2889 | -104.3916
18.2926 | -104.3994
18.2963 | -104.4072
18.3 | -104.4168
18.3037 | -104.4279
18.3074 | -104.4394
18.3111 | -104.4537
18.3148 | -104.4693
18.3185 | -104.4853
18.3222 | -104.5027
18.3259 | -104.5179
18.3296 | -104.5361
18.3333 | -104.5516
18.337 | -104.5693
18.3407 | -104.5875
18.3444 | -104.603
18.3481 | -104.6219
18.3519 | -104.6435
18.3556 | -104.6652
18.3593 | -104.6794
18.363 | -104.6945
18.3667 | -104.7114
18.3704 | -104.7279
18.3741 | -104.7413
18.3778 | -104.7562
18.3815 | -104.7684
18.3852 | -104.7781
18.3889 | -104.7878
18.3926 | -104.7997
18.3963 | -104.8087
18.4 | -104.8149
18.4037 | -104.8204
18.4074 | -104.8269
18.4111 | -104.8347
18.4148 | -104.8406
18.4185 | -104.8438
18.4222 | -104.8438
18.4259 | -104.8446
18.4296 | -104.8421
18.4333 | -104.8402
18.437 | -104.8381
18.4407 | -104.8338
18.4444 | -104.8294
18.4481 | -104.8259
18.4519 | -104.8253
18.4556 | -104.8248
18.4593 | -104.8238
18.463 | -104.8228
18.4667 | -104.8183
18.4704 | -104.8154
18.4741 | -104.816
18.4778 | -104.8153
18.4815 | -104.8124
18.4852 | -104.8109
18.4889 | -104.8122
18.4926 | -104.8113
18.4963 | -104.8125
18.5 | -104.8149
18.5037 | -104.8175
18.5074 | -104.8193
18.5111 | -104.8215
18.5148 | -104.8208
18.5185 | -104.819
18.5222 | -104.8201
18.5259 | -104.8211
18.5296 | -104.8227
18.5333 | -104.825
18.537 | -104.8243
18.5407 | -104.82
18.5444 | -104.8169
18.5481 | -104.8144
18.5519 | -104.8129
18.5556 | -104.8097
18.5593 | -104.8067
18.563 | -104.8048
18.5667 | -104.8012
18.5704 | -104.7952
18.5741 | -104.7888
18.5778 | -104.7823
18.5815 | -104.7761
18.5852 | -104.7692
18.5889 | -104.7641
18.5926 | -104.7537
18.5963 | -104.7437
18.6 | -104.7323
18.6037 | -104.722
18.6074 | -104.7139
18.6111 | -104.7057
18.6148 | -104.6925
18.6185 | -104.6787
18.6222 | -104.6687
18.6259 | -104.6592
18.6296 | -104.648
18.6333 | -104.6367
18.637 | -104.6231
18.6407 | -104.6089
18.6444 | -104.5985
18.6481 | -104.5875
18.6519 | -104.5721
18.6556 | -104.5527
18.6593 | -104.5353
18.663 | -104.5185
18.6667 | -104.5027
18.6704 | -104.4873
18.6741 | -104.4697
18.6778 | -104.4519
18.6815 | -104.4369
18.6852 | -104.4228
18.6889 | -104.4083
18.6926 | -104.3943
18.6963 | -104.3805
18.7 | -104.3727
18.7037 | -104.3618
18.7074 | -104.3535
18.7111 | -104.3439
18.7148 | -104.337
18.7185 | -104.3314
18.7222 | -104.3272
18.7259 | -104.3244
18.7296 | -104.324
18.7333 | -104.3258
18.737 | -104.3272
18.7407 | -104.3282
18.7444 | -104.334
18.7481 | -104.3441
18.7519 | -104.3555
18.7556 | -104.3644
18.7593 | -104.3755
18.763 | -104.3886
18.7667 | -104.4039
18.7704 | -104.4195
18.7741 | -104.4353
18.7778 | -104.4531
18.7815 | -104.4721
18.7852 | -104.4913
18.7889 | -104.5101
18.7926 | -104.5294
18.7963 | -104.5497
18.8 | -104.5685
18.8037 | -104.5889
18.8074 | -104.6089
18.8111 | -104.6271
18.8148 | -104.6468
18.8185 | -104.6652
18.8222 | -104.6795
18.8259 | -104.6922
18.8296 | -104.7062
18.8333 | -104.7204
18.837 | -104.7334
18.8407 | -104.7445
18.8444 | -104.7547
18.8481 | -104.7623
18.8519 | -104.7695
18.8556 | -104.7754
18.8593 | -104.7826
18.863 | -104.7872
18.8667 | -104.7855
18.8704 | -104.7822
18.8741 | -104.7788
18.8778 | -104.7733
18.8815 | -104.7696
18.8852 | -104.7638
18.8889 | -104.7538
18.8926 | -104.7446
18.8963 | -104.7356
18.9 | -104.7266
18.9037 | -104.7185
18.9074 | -104.7096
18.9111 | -104.7005
18.9148 | -104.6909
18.9185 | -104.6812
18.9222 | -104.6741
18.9259 | -104.667
18.9296 | -104.662
18.9333 | -104.6568
18.937 | -104.6504
18.9407 | -104.6473
18.9444 | -104.6445
18.9481 | -104.64
18.9519 | -104.6353
18.9556 | -104.6329
18.9593 | -104.6335
18.963 | -104.635
18.9667 | -104.6376
18.9704 | -104.6405
18.9741 | -104.6427
18.9778 | -104.6454
18.9815 | -104.6506
18.9852 | -104.6555
18.9889 | -104.6604
18.9926 | -104.6673
18.9963 | -104.6747
19 | -104.6825
19.0037 | -104.6889
19.0074 | -104.697
19.0111 | -104.7077
19.0148 | -104.7137
19.0185 | -104.7187
19.0222 | -104.7244
19.0259 | -104.7334
19.0296 | -104.7425
19.0333 | -104.7494
19.037 | -104.7546
19.0407 | -104.7581
19.0444 | -104.7616
19.0481 | -104.7668
19.0519 | -104.773
19.0556 | -104.781
19.0593 | -104.7844
19.063 | -104.7871
19.0667 | -104.7899
19.0704 | -104.7924
19.0741 | -104.794
19.0778 | -104.7967
19.0815 | -104.7988
19.0852 | -104.7987
19.0889 | -104.7975
19.0926 | -104.797
19.0963 | -104.7971
19.1 | -104.7962
19.1037 | -104.7939
19.1074 | -104.7911
19.1111 | -104.7906
19.1148 | -104.789
19.1185 | -104.785
19.1222 | -104.7826
19.1259 | -104.7791
19.1296 | -104.775
19.1333 | -104.7743
19.137 | -104.7783
19.1407 | -104.7785
19.1444 | -104.7766
19.1481 | -104.7762
19.1519 | -104.7763
19.1556 | -104.777
19.1593 | -104.7773
19.163 | -104.7748
19.1667 | -104.7743
19.1704 | -104.7756
19.1741 | -104.7772
19.1778 | -104.7798
19.1815 | -104.7823
19.1852 | -104.7863
19.1889 | -104.7909
19.1926 | -104.7937
19.1963 | -104.7955
19.2 | -104.796
19.2037 | -104.7974
19.2074 | -104.8014
19.2111 | -104.8053
19.2148 | -104.8092
19.2185 | -104.8141
19.2222 | -104.8218
19.2259 | -104.8283
19.2296 | -104.8361
19.2333 | -104.8445
19.237 | -104.8485
19.2407 | -104.8526
19.2444 | -104.8584
19.2481 | -104.8664
19.2519 | -104.8735
19.2556 | -104.8786
19.2593 | -104.8849
19.263 | -104.8911
19.2667 | -104.8941
19.2704 | -104.8985
19.2741 | -104.9013
19.2778 | -104.9048
19.2815 | -104.9083
19.2852 | -104.9092
19.2889 | -104.9131
19.2926 | -104.9162
19.2963 | -104.9176
19.3 | -104.9171
19.3037 | -104.9161
19.3074 | -104.9144
19.3111 | -104.911
19.3148 | -104.9089
19.3185 | -104.907
19.3222 | -104.9007
19.3259 | -104.8944
19.3296 | -104.8881
19.3333 | -104.8821
19.337 | -104.876
19.3407 | -104.8685
19.3444 | -104.862
19.3481 | -104.8542
19.3519 | -104.8441
19.3556 | -104.8357
19.3593 | -104.8266
19.363 | -104.8181
19.3667 | -104.8109
19.3704 | -104.803
19.3741 | -104.7951
19.3778 | -104.789
19.3815 | -104.786
19.3852 | -104.7812
19.3889 | -104.7762
19.3926 | -104.7714
19.3963 | -104.764
19.4 | -104.7601
19.4037 | -104.7583
19.4074 | -104.7542
19.4111 | -104.7496
19.4148 | -104.7474
19.4185 | -104.7469
19.4222 | -104.7473
19.4259 | -104.7487
19.4296 | -104.7508
19.4333 | -104.7544
19.437 | -104.7579
19.4407 | -104.7599
19.4444 | -104.7616
19.4481 | -104.7644
19.4519 | -104.7695
19.4556 | -104.7752
19.4593 | -104.7819
19.463 | -104.7903
19.4667 | -104.7985
19.4704 | -104.8046
19.4741 | -104.8113
19.4778 | -104.8166
19.4815 | -104.8218
19.4852 | -104.8258
19.4889 | -104.8297
19.4926 | -104.8321
19.4963 | -104.8317
19.5 | -104.8305
19.5037 | -104.83
19.5074 | -104.8287
19.5111 | -104.8267
19.5148 | -104.8258
19.5185 | -104.8228
19.5222 | -104.819
19.5259 | -104.8128
19.5296 | -104.8078
19.5333 | -104.8024
19.537 | -104.7954
19.5407 | -104.7897
19.5444 | -104.7849
19.5481 | -104.7795
19.5519 | -104.7738
19.5556 | -104.767
19.5593 | -104.7591
19.563 | -104.7509
19.5667 | -104.7435
19.5704 | -104.7357
19.5741 | -104.7289
19.5778 | -104.7236
19.5815 | -104.7185
19.5852 | -104.7137
19.5889 | -104.7096
19.5926 | -104.7043
19.5963 | -104.7012
19.6 | -104.7017
19.6037 | -104.7011
19.6074 | -104.7001
19.6111 | -104.702
19.6148 | -104.7049
19.6185 | -104.707
19.6222 | -104.7112
19.6259 | -104.7181
19.6296 | -104.7251
19.6333 | -104.7312
19.637 | -104.7385
19.6407 | -104.7464
19.6444 | -104.7557
19.6481 | -104.7646
19.6519 | -104.7727
19.6556 | -104.7829
19.6593 | -104.794
19.663 | -104.8058
19.6667 | -104.8173
19.6704 | -104.8304
19.6741 | -104.8454
19.6778 | -104.8565
19.6815 | -104.8655
19.6852 | -104.8762
19.6889 | -104.8874
19.6926 | -104.8971
19.6963 | -104.907
19.7 | -104.9169
19.7037 | -104.9241
19.7074 | -104.9286
19.7111 | -104.9318
19.7148 | -104.939
19.7185 | -104.9441
19.7222 | -104.9477
19.7259 | -104.9484
19.7296 | -104.9456
19.7333 | -104.9433
19.737 | -104.9391
19.7407 | -104.9339
19.7444 | -104.927
19.7481 | -104.9201
19.7519 | -104.9125
19.7556 | -104.9064
19.7593 | -104.8949
19.763 | -104.8837
19.7667 | -104.8693
19.7704 | -104.8583
19.7741 | -104.8483
19.7778 | -104.8367
19.7815 | -104.8255
19.7852 | -104.8129
19.7889 | -104.8009
19.7926 | -104.7877
19.7963 | -104.775
19.8 | -104.7631
19.8037 | -104.7497
19.8074 | -104.7372
19.8111 | -104.722
19.8148 | -104.7089
19.8185 | -104.6952
19.8222 | -104.6809
19.8259 | -104.6701
19.8296 | -104.6589
19.8333 | -104.6487
19.837 | -104.6377
19.8407 | -104.6261
19.8444 | -104.6165
19.8481 | -104.6085
19.8519 | -104.6009
19.8556 | -104.5928
19.8593 | -104.5794
19.863 | -104.5686
19.8667 | -104.5607
19.8704 | -104.5535
19.8741 | -104.5451
19.8778 | -104.5353
19.8815 | -104.5262
19.8852 | -104.5178
19.8889 | -104.5099
19.8926 | -104.501
19.8963 | -104.4925
19.9 | -104.4836
19.9037 | -104.4734
19.9074 | -104.4631
19.9111 | -104.4533
19.9148 | -104.4429
19.9185 | -104.4315
19.9222 | -104.4236
19.9259 | -104.4153
19.9296 | -104.4074
19.9333 | -104.4006
19.937 | -104.3935
19.9407 | -104.3849
19.9444 | -104.377
19.9481 | -104.37
19.9519 | -104.3631
19.9556 | -104.3562
19.9593 | -104.3494
19.963 | -104.3444
19.9667 | -104.3382
19.9704 | -104.3327
19.9741 | -104.3273
19.9778 | -104.3221
19.9815 | -104.3173
19.9852 | -104.3147
19.9889 | -104.315
19.9926 | -104.314
19.9963 | -104.3096
20 | -104.3081
20.0037 | -104.3072
20.0074 | -104.305
20.0111 | -104.3032
20.0148 | -104.3007
20.0185 | -104.2999
20.0222 | -104.2996
20.0259 | -104.2994
20.0296 | -104.2998
20.0333 | -104.3012
20.037 | -104.3013
20.0407 | -104.3043
20.0444 | -104.3068
20.0481 | -104.3074
20.0519 | -104.3098
20.0556 | -104.3128
20.0593 | -104.3161
20.063 | -104.3209
20.0667 | -104.3253
20.0704 | -104.3284
20.0741 | -104.3333
20.0778 | -104.3383
20.0815 | -104.3423
20.0852 | -104.3482
20.0889 | -104.3539
20.0926 | -104.3587
20.0963 | -104.3674
20.1 | -104.3754
20.1037 | -104.3841
20.1074 | -104.3926
20.1111 | -104.3987
20.1148 | -104.4073
20.1185 | -104.4171
20.1222 | -104.4269
20.1259 | -104.4382
20.1296 | -104.4479
20.1333 | -104.4559
20.137 | -104.4652
20.1407 | -104.4737
20.1444 | -104.4874
20.1481 | -104.5006
20.1519 | -104.5108
20.1556 | -104.5212
20.1593 | -104.5344
20.163 | -104.5474
20.1667 | -104.558
20.1704 | -104.5688
20.1741 | -104.5794
20.1778 | -104.5906
20.1815 | -104.6024
20.1852 | -104.615
20.1889 | -104.6273
20.1926 | -104.6389
20.1963 | -104.6505
20.2 | -104.6613
20.2037 | -104.6722
20.2074 | -104.6809
20.2111 | -104.6903
20.2148 | -104.6998
20.2185 | -104.7073
20.2222 | -104.7144
20.2259 | -104.7217
20.2296 | -104.7293
20.2333 | -104.7362
20.237 | -104.7448
20.2407 | -104.7523
20.2444 | -104.7583
20.2481 | -104.7659
20.2519 | -104.7733
20.2556 | -104.7789
20.2593 | -104.7863
20.263 | -104.792
20.2667 | -104.7951
20.2704 | -104.7978
20.2741 | -104.8018
20.2778 | -104.8069
20.2815 | -104.8126
20.2852 | -104.815
20.2889 | -104.8157
20.2926 | -104.819
20.2963 | -104.8228
20.3 | -104.8256
20.3037 | -104.8217
20.3074 | -104.8238
20.3111 | -104.8233
20.3148 | -104.8211
20.3185 | -104.8198
20.3222 | -104.8179
20.3259 | -104.8162
20.3296 | -104.8135
20.3333 | -104.8115
20.337 | -104.8096
20.3407 | -104.8034
20.3444 | -104.7962
20.3481 | -104.789
20.3519 | -104.7825
20.3556 | -104.7769
20.3593 | -104.7713
20.363 | -104.7668
20.3667 | -104.7642
20.3704 | -104.7586
20.3741 | -104.7522
20.3778 | -104.7471
20.3815 | -104.742
20.3852 | -104.7346
20.3889 | -104.7286
20.3926 | -104.7229
20.3963 | -104.7155
20.4 | -104.7098
20.4037 | -104.7024
20.4074 | -104.6949
20.4111 | -104.6893
20.4148 | -104.6826
20.4185 | -104.6765
20.4222 | -104.6701
20.4259 | -104.6633
20.4296 | -104.6582
20.4333 | -104.6504
20.437 | -104.6432
20.4407 | -104.6374
20.4444 | -104.6298
20.4481 | -104.6214
20.4519 | -104.6112
20.4556 | -104.6026
20.4593 | -104.593
20.463 | -104.582
20.4667 | -104.5723
20.4704 | -104.5602
20.4741 | -104.5493
20.4778 | -104.5389
20.4815 | -104.5278
20.4852 | -104.5169
20.4889 | -104.5072
20.4926 | -104.4986
20.4963 | -104.4874
20.5 | -104.4737
20.5037 | -104.4605
20.5074 | -104.448
20.5111 | -104.4331
20.5148 | -104.4188
20.5185 | -104.4049
20.5222 | -104.3922
20.5259 | -104.3786
20.5296 | -104.3644
20.5333 | -104.3509
20.537 | -104.3349
20.5407 | -104.3204
20.5444 | -104.3065
20.5481 | -104.2926
20.5519 | -104.2803
20.5556 | -104.268
20.5593 | -104.2563
20.563 | -104.2418
20.5667 | -104.2281
20.5704 | -104.2161
20.5741 | -104.2032
20.5778 | -104.1931
20.5815 | -104.1847
20.5852 | -104.1762
20.5889 | -104.1677
20.5926 | -104.1611
20.5963 | -104.155
20.6 | -104.1483
20.6037 | -104.1465
20.6074 | -104.1467
20.6111 | -104.1455
20.6148 | -104.1447
20.6185 | -104.1443
20.6222 | -104.1471
20.6259 | -104.1492
20.6296 | -104.1519
20.6333 | -104.1583
20.637 | -104.1659
20.6407 | -104.1733
20.6444 | -104.181
20.6481 | -104.188
20.6519 | -104.1962
20.6556 | -104.2029
20.6593 | -104.2106
20.663 | -104.2209
20.6667 | -104.2316
20.6704 | -104.243
20.6741 | -104.2564
20.6778 | -104.2696
20.6815 | -104.2831
20.6852 | -104.297
20.6889 | -104.312
20.6926 | -104.3273
20.6963 | -104.3423
20.7 | -104.3544
20.7037 | -104.3661
20.7074 | -104.3804
20.7111 | -104.3924
20.7148 | -104.406
20.7185 | -104.4204
20.7222 | -104.4327
20.7259 | -104.4444
20.7296 | -104.4543
20.7333 | -104.4626
20.737 | -104.4696
20.7407 | -104.4769
20.7444 | -104.4856
20.7481 | -104.4946
20.7519 | -104.5032
20.7556 | -104.5115
20.7593 | -104.5175
20.763 | -104.5216
20.7667 | -104.5223
20.7704 | -104.5214
20.7741 | -104.5208
20.7778 | -104.5172
20.7815 | -104.5123
20.7852 | -104.5061
20.7889 | -104.5007
20.7926 | -104.4961
20.7963 | -104.4917
20.8 | -104.4858
20.8037 | -104.4804
20.8074 | -104.4723
20.8111 | -104.4625
20.8148 | -104.4511
20.8185 | -104.4386
20.8222 | -104.4275
20.8259 | -104.4159
20.8296 | -104.403
20.8333 | -104.3898
20.837 | -104.378
20.8407 | -104.3684
20.8444 | -104.3587
20.8481 | -104.3471
20.8519 | -104.3349
20.8556 | -104.3227
20.8593 | -104.3133
20.863 | -104.306
20.8667 | -104.2969
20.8704 | -104.2888
20.8741 | -104.2797
20.8778 | -104.2703
20.8815 | -104.261
20.8852 | -104.254
20.8889 | -104.2478
20.8926 | -104.2408
20.8963 | -104.2355
20.9 | -104.2287
20.9037 | -104.2233
20.9074 | -104.219
20.9111 | -104.2153
20.9148 | -104.2147
20.9185 | -104.2147
20.9222 | -104.2143
20.9259 | -104.2153
20.9296 | -104.2153
20.9333 | -104.2167
20.937 | -104.2197
20.9407 | -104.2238
20.9444 | -104.2283
20.9481 | -104.2337
20.9519 | -104.2399
20.9556 | -104.247
20.9593 | -104.2574
20.963 | -104.268
20.9667 | -104.2772
20.9704 | -104.2858
20.9741 | -104.2962
20.9778 | -104.3063
20.9815 | -104.3182
20.9852 | -104.3314
20.9889 | -104.3448
20.9926 | -104.36
20.9963 | -104.3785
21 | -104.3964
21.0037 | -104.4124
21.0074 | -104.4293
21.0111 | -104.4441
21.0148 | -104.4605
21.0185 | -104.479
21.0222 | -104.4963
21.0259 | -104.5111
21.0296 | -104.53
21.0333 | -104.5511
21.037 | -104.5683
21.0407 | -104.5815
21.0444 | -104.5958
21.0481 | -104.611
21.0519 | -104.6259
21.0556 | -104.6397
21.0593 | -104.6537
21.063 | -104.6699
21.0667 | -104.685
21.0704 | -104.6988
21.0741 | -104.713
21.0778 | -104.7271
21.0815 | -104.7412
21.0852 | -104.7515
21.0889 | -104.7597
21.0926 | -104.765
21.0963 | -104.7724
21.1 | -104.7788
21.1037 | -104.7849
21.1074 | -104.7898
21.1111 | -104.7943
21.1148 | -104.7997
21.1185 | -104.805
21.1222 | -104.8112
21.1259 | -104.8159
21.1296 | -104.8199
21.1333 | -104.822
21.137 | -104.8211
21.1407 | -104.8184
21.1444 | -104.8144
21.1481 | -104.8129
21.1519 | -104.8143
21.1556 | -104.8143
21.1593 | -104.8118
21.163 | -104.8081
21.1667 | -104.8063
21.1704 | -104.8045
21.1741 | -104.7996
21.1778 | -104.7966
21.1815 | -104.795
21.1852 | -104.7923
21.1889 | -104.7899
21.1926 | -104.7885
21.1963 | -104.786
21.2 | -104.7848
21.2037 | -104.7816
21.2074 | -104.7785
21.2111 | -104.7771
21.2148 | -104.7737
21.2185 | -104.7707
21.2222 | -104.7715
21.2259 | -104.7713
21.2296 | -104.7718
21.2333 | -104.7702
21.237 | -104.7676
21.2407 | -104.7644
21.2444 | -104.7606
21.2481 | -104.7572
21.2519 | -104.7549
21.2556 | -104.7531
21.2593 | -104.7467
21.263 | -104.7404
21.2667 | -104.737
21.2704 | -104.733
21.2741 | -104.7292
21.2778 | -104.7243
21.2815 | -104.7165
21.2852 | -104.709
21.2889 | -104.7017
21.2926 | -104.6924
21.2963 | -104.6839
21.3 | -104.676
21.3037 | -104.6683
21.3074 | -104.6609
21.3111 | -104.6535
21.3148 | -104.644
21.3185 | -104.6341
21.3222 | -104.6248
21.3259 | -104.6154
21.3296 | -104.6032
21.3333 | -104.5917
21.337 | -104.5809
21.3407 | -104.5712
21.3444 | -104.5608
21.3481 | -104.548
21.3519 | -104.5342
21.3556 | -104.5208
21.3593 | -104.5092
21.363 | -104.4979
21.3667 | -104.4859
21.3704 | -104.4771
21.3741 | -104.4661
21.3778 | -104.4544
21.3815 | -104.4423
21.3852 | -104.4301
21.3889 | -104.4215
21.3926 | -104.4143
21.3963 | -104.4091
21.4 | -104.4028
21.4037 | -104.3973
21.4074 | -104.3934
21.4111 | -104.3888
21.4148 | -104.3838
21.4185 | -104.3777
21.4222 | -104.3707
21.4259 | -104.3649
21.4296 | -104.3593
21.4333 | -104.3561
21.437 | -104.3555
21.4407 | -104.3566
21.4444 | -104.3587
21.4481 | -104.3614
21.4519 | -104.363
21.4556 | -104.3658
21.4593 | -104.37
21.463 | -104.3749
21.4667 | -104.3781
21.4704 | -104.3817
21.4741 | -104.3882
21.4778 | -104.3959
21.4815 | -104.3998
21.4852 | -104.4045
21.4889 | -104.4107
21.4926 | -104.4158
21.4963 | -104.4218
21.5 | -104.4287
21.5037 | -104.4357
21.5074 | -104.4432
21.5111 | -104.4506
21.5148 | -104.4568
21.5185 | -104.4621
21.5222 | -104.4683
21.5259 | -104.473
21.5296 | -104.4771
21.5333 | -104.4822
21.537 | -104.4888
21.5407 | -104.4954
21.5444 | -104.5013
21.5481 | -104.5072
21.5519 | -104.5118
21.5556 | -104.5144
21.5593 | -104.5169
21.563 | -104.5228
21.5667 | -104.5261
21.5704 | -104.5279
21.5741 | -104.5292
21.5778 | -104.528
21.5815 | -104.5273
21.5852 | -104.5293
21.5889 | -104.5296
21.5926 | -104.5275
21.5963 | -104.5262
21.6 | -104.5228
21.6037 | -104.5182
21.6074 | -104.5149
21.6111 | -104.5113
21.6148 | -104.5073
21.6185 | -104.5042
21.6222 | -104.5004
21.6259 | -104.4981
21.6296 | -104.4958
21.6333 | -104.4909
21.637 | -104.4852
21.6407 | -104.4793
21.6444 | -104.472
21.6481 | -104.4645
21.6519 | -104.4592
21.6556 | -104.4532
21.6593 | -104.4464
21.663 | -104.4407
21.6667 | -104.4328
21.6704 | -104.4266
21.6741 | -104.4194
21.6778 | -104.4138
21.6815 | -104.4079
21.6852 | -104.4035
21.6889 | -104.3985
21.6926 | -104.3944
21.6963 | -104.3897
21.7 | -104.387
21.7037 | -104.3835
21.7074 | -104.3811
21.7111 | -104.3789
21.7148 | -104.3764
21.7185 | -104.3752
21.7222 | -104.3757
21.7259 | -104.3764
21.7296 | -104.379
21.7333 | -104.3814
21.737 | -104.3846
21.7407 | -104.3869
21.7444 | -104.3908
21.7481 | -104.3933
21.7519 | -104.3981
21.7556 | -104.4039
21.7593 | -104.4107
21.763 | -104.4157
21.7667 | -104.421
21.7704 | -104.4277
21.7741 | -104.4354
21.7778 | -104.4423
21.7815 | -104.4499
21.7852 | -104.4582
21.7889 | -104.4692
21.7926 | -104.4793
21.7963 | -104.4872
21.8 | -104.4949
21.8037 | -104.5031
21.8074 | -104.5097
21.8111 | -104.5151
21.8148 | -104.5218
21.8185 | -104.5291
21.8222 | -104.5354
21.8259 | -104.5439
21.8296 | -104.5506
21.8333 | -104.5533
21.837 | -104.5559
21.8407 | -104.5569
21.8444 | -104.5554
21.8481 | -104.5509
21.8519 | -104.5482
21.8556 | -104.547
21.8593 | -104.5453
21.863 | -104.5438
21.8667 | -104.5377
21.8704 | -104.5307
21.8741 | -104.5244
21.8778 | -104.5183
21.8815 | -104.5114
21.8852 | -104.5033
21.8889 | -104.4948
21.8926 | -104.4825
21.8963 | -104.4695
21.9 | -104.4578
21.9037 | -104.4452
21.9074 | -104.4322
21.9111 | -104.4217
21.9148 | -104.4104
21.9185 | -104.3965
21.9222 | -104.3818
21.9259 | -104.3674
21.9296 | -104.3517
21.9333 | -104.3373
21.937 | -104.3242
21.9407 | -104.3121
21.9444 | -104.3026
21.9481 | -104.2943
21.9519 | -104.2867
21.9556 | -104.2813
21.9593 | -104.2737
21.963 | -104.2665
21.9667 | -104.2613
21.9704 | -104.2562
21.9741 | -104.2515
21.9778 | -104.2475
21.9815 | -104.244
21.9852 | -104.2423
21.9889 | -104.2401
21.9926 | -104.2405
21.9963 | -104.2435
22 | -104.2473
22.0037 | -104.2505
22.0074 | -104.2529
22.0111 | -104.2535
22.0148 | -104.2561
22.0185 | -104.2595
22.0222 | -104.2644
22.0259 | -104.27
22.0296 | -104.2768
22.0333 | -104.283
22.037 | -104.2898
22.0407 | -104.2943
22.0444 | -104.3
22.0481 | -104.3063
22.0519 | -104.312
22.0556 | -104.3151
22.0593 | -104.3208
22.063 | -104.3258
22.0667 | -104.3303
22.0704 | -104.3345
22.0741 | -104.3385
22.0778 | -104.3422
22.0815 | -104.3457
22.0852 | -104.3491
22.0889 | -104.3522
22.0926 | -104.3553
22.0963 | -104.3592
22.1 | -104.3596
22.1037 | -104.3596
22.1074 | -104.3613
22.1111 | -104.3626
22.1148 | -104.3651
22.1185 | -104.3679
22.1222 | -104.3703
22.1259 | -104.3726
22.1296 | -104.3733
22.1333 | -104.3737
22.137 | -104.3718
22.1407 | -104.3702
22.1444 | -104.3734
22.1481 | -104.3759
22.1519 | -104.3786
22.1556 | -104.3795
22.1593 | -104.3813
22.163 | -104.3829
22.1667 | -104.384
22.1704 | -104.3836
22.1741 | -104.3828
22.1778 | -104.385
22.1815 | -104.386
22.1852 | -104.3854
22.1889 | -104.3851
22.1926 | -104.3852
22.1963 | -104.3871
22.2 | -104.388
22.2037 | -104.3899
22.2074 | -104.392
22.2111 | -104.3914
22.2148 | -104.3903
22.2185 | -104.3889
22.2222 | -104.3892
22.2259 | -104.39
22.2296 | -104.3913
22.2333 | -104.3928
22.237 | -104.3939
22.2407 | -104.395
22.2444 | -104.3966
22.2481 | -104.3971
22.2519 | -104.3972
22.2556 | -104.3967
22.2593 | -104.3958
22.263 | -104.3955
22.2667 | -104.3943
22.2704 | -104.394
22.2741 | -104.394
22.2778 | -104.3947
22.2815 | -104.3948
22.2852 | -104.3957
22.2889 | -104.3958
22.2926 | -104.3967
22.2963 | -104.3962
22.3 | -104.3978
22.3037 | -104.3997
22.3074 | -104.4007
22.3111 | -104.4005
22.3148 | -104.4015
22.3185 | -104.4008
22.3222 | -104.4028
22.3259 | -104.4061
22.3296 | -104.4072
22.3333 | -104.4079
22.337 | -104.4076
22.3407 | -104.4046
22.3444 | -104.4051
22.3481 | -104.4053
22.3519 | -104.4023
22.3556 | -104.3996
22.3593 | -104.3964
22.363 | -104.3933
22.3667 | -104.3901
22.3704 | -104.3866
22.3741 | -104.3826
22.3778 | -104.3782
22.3815 | -104.3723
22.3852 | -104.3657
22.3889 | -104.36
22.3926 | -104.3532
22.3963 | -104.345
22.4 | -104.336
22.4037 | -104.3259
22.4074 | -104.3147
22.4111 | -104.3048
22.4148 | -104.2941
22.4185 | -104.2814
22.4222 | -104.2692
22.4259 | -104.2591
22.4296 | -104.2463
22.4333 | -104.2338
22.437 | -104.2221
22.4407 | -104.2071
22.4444 | -104.1899
22.4481 | -104.1761
22.4519 | -104.164
22.4556 | -104.1525
22.4593 | -104.1394
22.463 | -104.1231
22.4667 | -104.1082
22.4704 | -104.0947
22.4741 | -104.0817
22.4778 | -104.0693
22.4815 | -104.0558
22.4852 | -104.0417
22.4889 | -104.028
22.4926 | -104.0155
22.4963 | -104.0037
22.5 | -103.9923
22.5037 | -103.9812
22.5074 | -103.9689
22.5111 | -103.9566
22.5148 | -103.9423
22.5185 | -103.9299
22.5222 | -103.9192
22.5259 | -103.9087
22.5296 | -103.8989
22.5333 | -103.8888
22.537 | -103.8789
22.5407 | -103.8678
22.5444 | -103.8575
22.5481 | -103.8485
22.5519 | -103.8404
22.5556 | -103.833
22.5593 | -103.8255
22.563 | -103.8183
22.5667 | -103.8137
22.5704 | -103.8079
22.5741 | -103.8031
22.5778 | -103.7992
22.5815 | -103.7945
22.5852 | -103.7918
22.5889 | -103.7889
22.5926 | -103.7869
22.5963 | -103.7856
22.6 | -103.7868
22.6037 | -103.7867
22.6074 | -103.7858
22.6111 | -103.7851
22.6148 | -103.7874
22.6185 | -103.7899
22.6222 | -103.7908
22.6259 | -103.7935
22.6296 | -103.7968
22.6333 | -103.8001
22.637 | -103.8062
22.6407 | -103.8136
22.6444 | -103.8205
22.6481 | -103.8263
22.6519 | -103.8324
22.6556 | -103.84
22.6593 | -103.85
22.663 | -103.857
22.6667 | -103.8638
22.6704 | -103.8735
22.6741 | -103.8829
22.6778 | -103.8917
22.6815 | -103.9021
22.6852 | -103.9105
22.6889 | -103.9186
22.6926 | -103.926
22.6963 | -103.9337
22.7 | -103.9403
22.7037 | -103.9476
22.7074 | -103.9549
22.7111 | -103.962
22.7148 | -103.9692
22.7185 | -103.9786
22.7222 | -103.9881
22.7259 | -103.9965
22.7296 | -104.0022
22.7333 | -104.0095
22.737 | -104.0181
22.7407 | -104.0251
22.7444 | -104.0303
22.7481 | -104.0348
22.7519 | -104.039
22.7556 | -104.0433
22.7593 | -104.0477
22.763 | -104.0525
22.7667 | -104.0588
22.7704 | -104.0613
22.7741 | -104.0625
22.7778 | -104.0627
22.7815 | -104.0623
22.7852 | -104.0635
22.7889 | -104.0625
22.7926 | -104.0591
22.7963 | -104.0572
22.8 | -104.0543
22.8037 | -104.0518
22.8074 | -104.0483
22.8111 | -104.0463
22.8148 | -104.0429
22.8185 | -104.0383
22.8222 | -104.0326
22.8259 | -104.0288
22.8296 | -104.0247
22.8333 | -104.0204
22.837 | -104.0148
22.8407 | -104.0041
22.8444 | -103.9927
22.8481 | -103.9806
22.8519 | -103.9699
22.8556 | -103.96
22.8593 | -103.949
22.863 | -103.9391
22.8667 | -103.9291
22.8704 | -103.9182
22.8741 | -103.907
22.8778 | -103.8982
22.8815 | -103.8879
22.8852 | -103.8777
22.8889 | -103.8657
22.8926 | -103.8524
22.8963 | -103.8408
22.9 | -103.8305
22.9037 | -103.8201
22.9074 | -103.8097
22.9111 | -103.8003
22.9148 | -103.7908
22.9185 | -103.7806
22.9222 | -103.7707
22.9259 | -103.7635
22.9296 | -103.7563
22.9333 | -103.7477
22.937 | -103.7386
22.9407 | -103.7305
22.9444 | -103.7253
22.9481 | -103.7211
22.9519 | -103.7172
22.9556 | -103.713
22.9593 | -103.7082
22.963 | -103.7035
22.9667 | -103.6992
22.9704 | -103.698
22.9741 | -103.7003
22.9778 | -103.699
22.9815 | -103.6976
22.9852 | -103.6963
22.9889 | -103.6962
22.9926 | -103.6982
22.9963 | -103.7009
23 | -103.7044
23.0037 | -103.7065
23.0074 | -103.7077
23.0111 | -103.7114
23.0148 | -103.7162
23.0185 | -103.7199
23.0222 | -103.7243
23.0259 | -103.7315
23.0296 | -103.7373
23.0333 | -103.7408
23.037 | -103.7449
23.0407 | -103.7491
23.0444 | -103.752
23.0481 | -103.7557
23.0519 | -103.7584
23.0556 | -103.7614
23.0593 | -103.7651
23.063 | -103.7692
23.0667 | -103.7731
23.0704 | -103.7767
23.0741 | -103.7815
23.0778 | -103.7833
23.0815 | -103.7846
23.0852 | -103.7866
23.0889 | -103.7894
23.0926 | -103.7934
23.0963 | -103.7967
23.1 | -103.797
23.1037 | -103.798
23.1074 | -103.801
23.1111 | -103.8029
23.1148 | -103.804
23.1185 | -103.805
23.1222 | -103.804
23.1259 | -103.8024
23.1296 | -103.7987
23.1333 | -103.7951
23.137 | -103.7914
23.1407 | -103.7877
23.1444 | -103.7831
23.1481 | -103.7789
23.1519 | -103.7743
23.1556 | -103.7689
23.1593 | -103.7629
23.163 | -103.7594
23.1667 | -103.757
23.1704 | -103.7527
23.1741 | -103.7474
23.1778 | -103.7437
23.1815 | -103.7383
23.1852 | -103.7314
23.1889 | -103.7228
23.1926 | -103.7148
23.1963 | -103.7061
23.2 | -103.6968
23.2037 | -103.6871
23.2074 | -103.6804
23.2111 | -103.6737
23.2148 | -103.6662
23.2185 | -103.6605
23.2222 | -103.6547
23.2259 | -103.6464
23.2296 | -103.6386
23.2333 | -103.6328
23.237 | -103.626
23.2407 | -103.6197
23.2444 | -103.6145
23.2481 | -103.6081
23.2519 | -103.6018
23.2556 | -103.5977
23.2593 | -103.5963
23.263 | -103.593
23.2667 | -103.587
23.2704 | -103.5804
23.2741 | -103.5748
23.2778 | -103.5694
23.2815 | -103.5656
23.2852 | -103.564
23.2889 | -103.5617
23.2926 | -103.562
23.2963 | -103.5615
23.3 | -103.5615
23.3037 | -103.5622
23.3074 | -103.5628
23.3111 | -103.5649
23.3148 | -103.5651
23.3185 | -103.5646
23.3222 | -103.5643
23.3259 | -103.5643
23.3296 | -103.5645
23.3333 | -103.5662
23.337 | -103.5679
23.3407 | -103.57
23.3444 | -103.5715
23.3481 | -103.5727
23.3519 | -103.5736
23.3556 | -103.5748
23.3593 | -103.5767
23.363 | -103.5789
23.3667 | -103.5789
23.3704 | -103.5787
23.3741 | -103.5779
23.3778 | -103.5762
23.3815 | -103.5753
23.3852 | -103.5739
23.3889 | -103.5701
23.3926 | -103.5648
23.3963 | -103.5608
23.4 | -103.5564
23.4037 | -103.5519
23.4074 | -103.5484
23.4111 | -103.5436
23.4148 | -103.5361
23.4185 | -103.5276
23.4222 | -103.5191
23.4259 | -103.5117
23.4296 | -103.5054
23.4333 | -103.4991
23.437 | -103.4907
23.4407 | -103.4827
23.4444 | -103.4739
23.4481 | -103.4665
23.4519 | -103.4612
23.4556 | -103.4572
23.4593 | -103.4533
23.463 | -103.4496
23.4667 | -103.4469
23.4704 | -103.4455
23.4741 | -103.4448
23.4778 | -103.4429
23.4815 | -103.4415
23.4852 | -103.4399
23.4889 | -103.4405
23.4926 | -103.4402
23.4963 | -103.4418
23.5 | -103.4425
23.5037 | -103.4441
23.5074 | -103.4525
23.5111 | -103.4535
23.5148 | -103.4563
23.5185 | -103.461
23.5222 | -103.4669
23.5259 | -103.4711
23.5296 | -103.476
23.5333 | -103.4811
23.537 | -103.4853
23.5407 | -103.4902
23.5444 | -103.4944
23.5481 | -103.5001
23.5519 | -103.5075
23.5556 | -103.5136
23.5593 | -103.5223
23.563 | -103.5313
23.5667 | -103.5387
23.5704 | -103.5473
23.5741 | -103.5559
23.5778 | -103.5651
23.5815 | -103.5755
23.5852 | -103.5852
23.5889 | -103.5942
23.5926 | -103.6039
23.5963 | -103.6152
23.6 | -103.6287
23.6037 | -103.6418
23.6074 | -103.6543
23.6111 | -103.6684
23.6148 | -103.6816
23.6185 | -103.6941
23.6222 | -103.7084
23.6259 | -103.7214
23.6296 | -103.7349
23.6333 | -103.7485
23.637 | -103.7633
23.6407 | -103.7792
23.6444 | -103.794
23.6481 | -103.811
23.6519 | -103.8278
23.6556 | -103.8446
23.6593 | -103.8615
23.663 | -103.8786
23.6667 | -103.8968
23.6704 | -103.9139
23.6741 | -103.9304
23.6778 | -103.9487
23.6815 | -103.9692
23.6852 | -103.9892
23.6889 | -104.0078
23.6926 | -104.0259
23.6963 | -104.0437
23.7 | -104.0612
23.7037 | -104.0788
23.7074 | -104.0964
23.7111 | -104.1132
23.7148 | -104.1293
23.7185 | -104.146
23.7222 | -104.1628
23.7259 | -104.1801
23.7296 | -104.1951
23.7333 | -104.2091
23.737 | -104.2219
23.7407 | -104.2342
23.7444 | -104.2461
23.7481 | -104.258
23.7519 | -104.2698
23.7556 | -104.2797
23.7593 | -104.287
23.763 | -104.2929
23.7667 | -104.2972
23.7704 | -104.3025
23.7741 | -104.305
23.7778 | -104.3085
23.7815 | -104.3073
23.7852 | -104.3044
23.7889 | -104.3001
23.7926 | -104.2942
23.7963 | -104.2885
23.8 | -104.2805
23.8037 | -104.2714
23.8074 | -104.2625
23.8111 | -104.2546
23.8148 | -104.2431
23.8185 | -104.2341
23.8222 | -104.2219
23.8259 | -104.2097
23.8296 | -104.1968
23.8333 | -104.1836
23.837 | -104.1709
23.8407 | -104.1571
23.8444 | -104.1442
23.8481 | -104.131
23.8519 | -104.1192
23.8556 | -104.1082
23.8593 | -104.0961
23.863 | -104.0862
23.8667 | -104.0785
23.8704 | -104.0727
23.8741 | -104.0674
23.8778 | -104.0595
23.8815 | -104.0531
23.8852 | -104.0485
23.8889 | -104.0443
23.8926 | -104.0413
23.8963 | -104.0375
23.9 | -104.0357
23.9037 | -104.0363
23.9074 | -104.0372
23.9111 | -104.0407
23.9148 | -104.0447
23.9185 | -104.0501
23.9222 | -104.0577
23.9259 | -104.0655
23.9296 | -104.0731
23.9333 | -104.0835
23.937 | -104.0935
23.9407 | -104.1042
23.9444 | -104.1152
23.9481 | -104.1279
23.9519 | -104.1423
23.9556 | -104.1565
23.9593 | -104.1703
23.963 | -104.1862
23.9667 | -104.2026
23.9704 | -104.2185
23.9741 | -104.2351
23.9778 | -104.2518
23.9815 | -104.2694
23.9852 | -104.2875
23.9889 | -104.3049
23.9926 | -104.3221
23.9963 | -104.3393
24 | -104.3575
24.0037 | -104.3752
24.0074 | -104.3935
24.0111 | -104.4122
24.0148 | -104.4293
24.0185 | -104.4463
24.0222 | -104.4617
24.0259 | -104.4791
24.0296 | -104.4963
24.0333 | -104.5143
24.037 | -104.5336
24.0407 | -104.552
24.0444 | -104.5657
24.0481 | -104.5829
24.0519 | -104.6002
24.0556 | -104.6168
24.0593 | -104.6321
24.063 | -104.6475
24.0667 | -104.6615
24.0704 | -104.6754
24.0741 | -104.6897
24.0778 | -104.7032
24.0815 | -104.7165
24.0852 | -104.7304
24.0889 | -104.7442
24.0926 | -104.7574
24.0963 | -104.7692
24.1 | -104.7836
24.1037 | -104.7968
24.1074 | -104.8079
24.1111 | -104.8182
24.1148 | -104.8291
24.1185 | -104.8398
24.1222 | -104.852
24.1259 | -104.8617
24.1296 | -104.8716
24.1333 | -104.8828
24.137 | -104.8932
24.1407 | -104.902
24.1444 | -104.9084
24.1481 | -104.9135
24.1519 | -104.9185
24.1556 | -104.9275
24.1593 | -104.9343
24.163 | -104.9403
24.1667 | -104.9447
24.1704 | -104.9487
24.1741 | -104.9546
24.1778 | -104.9587
24.1815 | -104.9615
24.1852 | -104.9648
24.1889 | -104.967
24.1926 | -104.9668
24.1963 | -104.9633
24.2 | -104.9622
24.2037 | -104.962
24.2074 | -104.9609
24.2111 | -104.9579
24.2148 | -104.9535
24.2185 | -104.9498
24.2222 | -104.9458
24.2259 | -104.9404
24.2296 | -104.9353
24.2333 | -104.9297
24.237 | -104.9238
24.2407 | -104.9163
24.2444 | -104.9072
24.2481 | -104.898
24.2519 | -104.8876
24.2556 | -104.8767
24.2593 | -104.8651
24.263 | -104.8547
24.2667 | -104.8434
24.2704 | -104.8312
24.2741 | -104.8167
24.2778 | -104.8031
24.2815 | -104.7892
24.2852 | -104.7759
24.2889 | -104.7606
24.2926 | -104.7432
24.2963 | -104.7266
24.3 | -104.7091
24.3037 | -104.6896
24.3074 | -104.6675
24.3111 | -104.6435
24.3148 | -104.6208
24.3185 | -104.6017
24.3222 | -104.5841
24.3259 | -104.5651
24.3296 | -104.545
24.3333 | -104.5237
24.337 | -104.5028
24.3407 | -104.4822
24.3444 | -104.4595
24.3481 | -104.4365
24.3519 | -104.4137
24.3556 | -104.3909
24.3593 | -104.3678
24.363 | -104.3451
24.3667 | -104.3207
24.3704 | -104.2966
24.3741 | -104.2712
24.3778 | -104.2443
24.3815 | -104.2167
24.3852 | -104.1898
24.3889 | -104.1643
24.3926 | -104.1391
24.3963 | -104.115
24.4 | -104.0908
24.4037 | -104.0659
24.4074 | -104.0414
24.4111 | -104.0192
24.4148 | -103.9971
24.4185 | -103.9751
24.4222 | -103.9551
24.4259 | -103.9371
24.4296 | -103.9196
24.4333 | -103.9033
24.437 | -103.8875
24.4407 | -103.8729
24.4444 | -103.859
24.4481 | -103.8481
24.4519 | -103.8403
24.4556 | -103.8327
24.4593 | -103.827
24.463 | -103.8237
24.4667 | -103.8203
24.4704 | -103.8187
24.4741 | -103.8188
24.4778 | -103.8207
24.4815 | -103.8242
24.4852 | -103.8309
24.4889 | -103.8407
24.4926 | -103.852
24.4963 | -103.863
24.5 | -103.874
24.5037 | -103.8877
24.5074 | -103.9024
24.5111 | -103.9187
24.5148 | -103.9344
24.5185 | -103.9511
24.5222 | -103.9689
24.5259 | -103.9859
24.5296 | -104.0043
24.5333 | -104.024
24.537 | -104.0438
24.5407 | -104.0661
24.5444 | -104.0884
24.5481 | -104.108
24.5519 | -104.1253
24.5556 | -104.1437
24.5593 | -104.164
24.563 | -104.1849
24.5667 | -104.2038
24.5704 | -104.2169
24.5741 | -104.2221
24.5778 | -104.2061
24.5815 | -104.2
24.5852 | -104.1897
24.5889 | -104.1654
24.5926 | -104.1272
24.5963 | -104.0874
24.6 | -104.049
24.6037 | -104.0061
24.6074 | -103.9561
24.6111 | -103.9059
24.6148 | -103.8594
24.6185 | -103.8101
24.6222 | -103.7574
24.6259 | -103.7
24.6296 | -103.6394
24.6333 | -103.5748
24.637 | -103.5078
24.6407 | -103.4386
24.6444 | -103.3652
24.6481 | -103.2892
24.6519 | -103.2102
24.6556 | -103.1289
24.6593 | -103.0473
24.663 | -102.9617
24.6667 | -102.8714
24.6704 | -102.7795
24.6741 | -102.6887
24.6778 | -102.5976
24.6815 | -102.5061
24.6852 | -102.4144
24.6889 | -102.3231
24.6926 | -102.2341
24.6963 | -102.1469
24.7 | -102.0614
24.7037 | -101.9784
24.7074 | -101.8975
24.7111 | -101.8201
24.7148 | -101.7459
24.7185 | -101.6748
24.7222 | -101.6081
24.7259 | -101.5459
24.7296 | -101.4882
24.7333 | -101.4349
24.737 | -101.3846
24.7407 | -101.3389
24.7444 | -101.2982
24.7481 | -101.2637
24.7519 | -101.2348
24.7556 | -101.2119
24.7593 | -101.1936
24.763 | -101.1808
24.7667 | -101.1741
24.7704 | -101.1725
24.7741 | -101.1755
24.7778 | -101.1843
24.7815 | -101.1993
24.7852 | -101.2206
24.7889 | -101.2469
24.7926 | -101.2784
24.7963 | -101.3154
24.8 | -101.3574
24.8037 | -101.4047
24.8074 | -101.4567
24.8111 | -101.5134
24.8148 | -101.5753
24.8185 | -101.6417
24.8222 | -101.7109
24.8259 | -101.7822
24.8296 | -101.8581
24.8333 | -101.9372
24.837 | -102.019
24.8407 | -102.1034
24.8444 | -102.1913
24.8481 | -102.2809
24.8519 | -102.3707
24.8556 | -102.4622
24.8593 | -102.5544
24.863 | -102.6478
24.8667 | -102.7405
24.8704 | -102.8334
24.8741 | -102.9255
24.8778 | -103.0149
24.8815 | -103.1037
24.8852 | -103.1915
24.8889 | -103.2771
24.8926 | -103.3594
24.8963 | -103.4399
24.9 | -103.5182
24.9037 | -103.5921
24.9074 | -103.6622
24.9111 | -103.7282
24.9148 | -103.7915
24.9185 | -103.8546
24.9222 | -103.918
24.9259 | -103.975
24.9296 | -104.0322
24.9333 | -104.0876
24.937 | -104.1391
24.9407 | -104.196
24.9444 | -104.2568
24.9481 | -104.3058
24.9519 | -104.3423
24.9556 | -104.3751
24.9593 | -104.4047
24.963 | -104.4228
24.9667 | -104.4306
24.9704 | -104.4369
24.9741 | -104.4407
24.9778 | -104.4433
24.9815 | -104.4451
24.9852 | -104.4453
24.9889 | -104.442
24.9926 | -104.4399
24.9963 | -104.4388
25 | -104.4355
25.0037 | -104.4333
25.0074 | -104.4323
25.0111 | -104.4298
25.0148 | -104.4285
25.0185 | -104.428
25.0222 | -104.4282
25.0259 | -104.4279
25.0296 | -104.431
25.0333 | -104.4348
25.037 | -104.4389
25.0407 | -104.4428
25.0444 | -104.4474
25.0481 | -104.4526
25.0519 | -104.4567
25.0556 | -104.4599
25.0593 | -104.4662
25.063 | -104.4737
25.0667 | -104.4797
25.0704 | -104.4867
25.0741 | -104.4944
25.0778 | -104.5014
25.0815 | -104.5082
25.0852 | -104.5135
25.0889 | -104.5204
25.0926 | -104.5293
25.0963 | -104.5377
25.1 | -104.5424
25.1037 | -104.5502
25.1074 | -104.5566
25.1111 | -104.5654
25.1148 | -104.5709
25.1185 | -104.578
25.1222 | -104.5853
25.1259 | -104.5929
25.1296 | -104.6
25.1333 | -104.605
25.137 | -104.6093
25.1407 | -104.6136
25.1444 | -104.6179
25.1481 | -104.6228
25.1519 | -104.6271
25.1556 | -104.6292
25.1593 | -104.6302
25.163 | -104.632
25.1667 | -104.6327
25.1704 | -104.6338
25.1741 | -104.6358
25.1778 | -104.6382
25.1815 | -104.642
25.1852 | -104.6427
25.1889 | -104.6411
25.1926 | -104.6373
25.1963 | -104.6342
25.2 | -104.6303
25.2037 | -104.6264
25.2074 | -104.6224
25.2111 | -104.6192
25.2148 | -104.6157
25.2185 | -104.6108
25.2222 | -104.6051
25.2259 | -104.5982
25.2296 | -104.5929
25.2333 | -104.5878
25.237 | -104.5815
25.2407 | -104.5736
25.2444 | -104.5653
25.2481 | -104.5575
25.2519 | -104.5515
25.2556 | -104.5462
25.2593 | -104.5416
25.263 | -104.5361
25.2667 | -104.5307
25.2704 | -104.5258
25.2741 | -104.521
25.2778 | -104.5153
25.2815 | -104.5122
25.2852 | -104.5077
25.2889 | -104.5035
25.2926 | -104.5009
25.2963 | -104.498
25.3 | -104.4961
25.3037 | -104.4962
25.3074 | -104.4969
25.3111 | -104.4963
25.3148 | -104.4962
25.3185 | -104.4972
25.3222 | -104.4994
25.3259 | -104.5
25.3296 | -104.5003
25.3333 | -104.5012
25.337 | -104.5025
25.3407 | -104.5033
25.3444 | -104.506
25.3481 | -104.5083
25.3519 | -104.509
25.3556 | -104.5111
25.3593 | -104.5123
25.363 | -104.5157
25.3667 | -104.519
25.3704 | -104.5213
25.3741 | -104.5224
25.3778 | -104.5234
25.3815 | -104.5252
25.3852 | -104.5253
25.3889 | -104.5245
25.3926 | -104.5231
25.3963 | -104.5233
25.4 | -104.5218
25.4037 | -104.5187
25.4074 | -104.5168
25.4111 | -104.5154
25.4148 | -104.5117
25.4185 | -104.5072
25.4222 | -104.5055
25.4259 | -104.5026
25.4296 | -104.4977
25.4333 | -104.494
25.437 | -104.4888
25.4407 | -104.4833
25.4444 | -104.479
25.4481 | -104.4716
25.4519 | -104.4637
25.4556 | -104.458
25.4593 | -104.45
25.463 | -104.4423
25.4667 | -104.4339
25.4704 | -104.4258
25.4741 | -104.418
25.4778 | -104.4104
25.4815 | -104.4032
25.4852 | -104.3957
25.4889 | -104.389
25.4926 | -104.3832
25.4963 | -104.3751
25.5 | -104.3645
25.5037 | -104.3566
25.5074 | -104.3476
25.5111 | -104.3393
25.5148 | -104.3332
25.5185 | -104.326
25.5222 | -104.3191
25.5259 | -104.3121
25.5296 | -104.3034
25.5333 | -104.295
25.537 | -104.2892
25.5407 | -104.284
25.5444 | -104.2796
25.5481 | -104.2767
25.5519 | -104.2734
25.5556 | -104.2702
25.5593 | -104.2677
25.563 | -104.265
25.5667 | -104.2624
25.5704 | -104.259
25.5741 | -104.2566
25.5778 | -104.2554
25.5815 | -104.2552
25.5852 | -104.2555
25.5889 | -104.2556
25.5926 | -104.2555
25.5963 | -104.2546
25.6 | -104.2559
25.6037 | -104.2573
25.6074 | -104.2588
25.6111 | -104.2591
25.6148 | -104.2597
25.6185 | -104.2607
25.6222 | -104.2629
25.6259 | -104.2656
25.6296 | -104.2701
25.6333 | -104.2738
25.637 | -104.277
25.6407 | -104.279
25.6444 | -104.2802
25.6481 | -104.2817
25.6519 | -104.2836
25.6556 | -104.286
25.6593 | -104.2879
25.663 | -104.2887
25.6667 | -104.2889
25.6704 | -104.2881
25.6741 | -104.2879
25.6778 | -104.2886
25.6815 | -104.2892
25.6852 | -104.2894
25.6889 | -104.2899
25.6926 | -104.2899
25.6963 | -104.2898
25.7 | -104.2895
25.7037 | -104.2891
25.7074 | -104.2882
25.7111 | -104.2865
25.7148 | -104.286
25.7185 | -104.2849
25.7222 | -104.2859
25.7259 | -104.2852
25.7296 | -104.2843
25.7333 | -104.2818
25.737 | -104.2792
25.7407 | -104.2754
25.7444 | -104.2728
25.7481 | -104.2717
25.7519 | -104.2725
25.7556 | -104.2736
25.7593 | -104.2724
25.763 | -104.2719
25.7667 | -104.2719
25.7704 | -104.2711
25.7741 | -104.2707
25.7778 | -104.2712
25.7815 | -104.2702
25.7852 | -104.2687
25.7889 | -104.2688
25.7926 | -104.2686
25.7963 | -104.2659
25.8 | -104.2655
25.8037 | -104.2662
25.8074 | -104.266
25.8111 | -104.2669
25.8148 | -104.2682
25.8185 | -104.2655
25.8222 | -104.2638
25.8259 | -104.2635
25.8296 | -104.2636
25.8333 | -104.2631
25.837 | -104.2637
25.8407 | -104.2638
25.8444 | -104.2636
25.8481 | -104.2622
25.8519 | -104.2632
25.8556 | -104.2642
25.8593 | -104.2654
25.863 | -104.267
25.8667 | -104.2666
25.8704 | -104.2667
25.8741 | -104.2661
25.8778 | -104.2666
25.8815 | -104.2692
25.8852 | -104.2686
25.8889 | -104.2686
25.8926 | -104.2706
25.8963 | -104.2698
25.9 | -104.2681
25.9037 | -104.2675
25.9074 | -104.2655
25.9111 | -104.2642
25.9148 | -104.2662
25.9185 | -104.2668
25.9222 | -104.2678
25.9259 | -104.269
25.9296 | -104.2698
25.9333 | -104.2721
25.937 | -104.2746
25.9407 | -104.2771
25.9444 | -104.2783
25.9481 | -104.2797
25.9519 | -104.2828
25.9556 | -104.2863
25.9593 | -104.2916
25.963 | -104.2941
25.9667 | -104.298
25.9704 | -104.3013
25.9741 | -104.3049
25.9778 | -104.3094
25.9815 | -104.3155
25.9852 | -104.3208
25.9889 | -104.3254
25.9926 | -104.3293
25.9963 | -104.3353
26 | -104.3424
26.0037 | -104.3484
26.0074 | -104.3554
26.0111 | -104.3634
26.0148 | -104.3719
26.0185 | -104.3795
26.0222 | -104.386
26.0259 | -104.3937
26.0296 | -104.401
26.0333 | -104.4102
26.037 | -104.4195
26.0407 | -104.4287
26.0444 | -104.4367
26.0481 | -104.4445
26.0519 | -104.4519
26.0556 | -104.4603
26.0593 | -104.4698
26.063 | -104.4788
26.0667 | -104.4874
26.0704 | -104.4971
26.0741 | -104.5061
26.0778 | -104.515
26.0815 | -104.5237
26.0852 | -104.5314
26.0889 | -104.5379
26.0926 | -104.5456
26.0963 | -104.5535
26.1 | -104.5627
26.1037 | -104.5724
26.1074 | -104.5806
26.1111 | -104.5875
26.1148 | -104.5947
26.1185 | -104.6004
26.1222 | -104.6056
26.1259 | -104.612
26.1296 | -104.6186
26.1333 | -104.6249
26.137 | -104.6313
26.1407 | -104.637
26.1444 | -104.6398
26.1481 | -104.6446
26.1519 | -104.6491
26.1556 | -104.6525
26.1593 | -104.657
26.163 | -104.6626
26.1667 | -104.6655
26.1704 | -104.6685
26.1741 | -104.6703
26.1778 | -104.6717
26.1815 | -104.6734
26.1852 | -104.6757
26.1889 | -104.6808
26.1926 | -104.6831
26.1963 | -104.6843
26.2 | -104.6852
26.2037 | -104.6862
26.2074 | -104.6867
26.2111 | -104.6858
26.2148 | -104.6864
26.2185 | -104.6872
26.2222 | -104.6873
26.2259 | -104.6875
26.2296 | -104.6902
26.2333 | -104.691
26.237 | -104.6907
26.2407 | -104.6906
26.2444 | -104.6912
26.2481 | -104.6896
26.2519 | -104.6886
26.2556 | -104.69
26.2593 | -104.6895
26.263 | -104.6884
26.2667 | -104.6886
26.2704 | -104.6875
26.2741 | -104.6861
26.2778 | -104.685
26.2815 | -104.6828
26.2852 | -104.681
26.2889 | -104.6793
26.2926 | -104.6759
26.2963 | -104.6723
26.3 | -104.669
26.3037 | -104.6644
26.3074 | -104.6592
26.3111 | -104.6535
26.3148 | -104.6489
26.3185 | -104.6434
26.3222 | -104.6366
26.3259 | -104.6309
26.3296 | -104.6239
26.3333 | -104.6194
26.337 | -104.6122
26.3407 | -104.6039
26.3444 | -104.5973
26.3481 | -104.5901
26.3519 | -104.5817
26.3556 | -104.5748
26.3593 | -104.5689
26.363 | -104.5611
26.3667 | -104.5529
26.3704 | -104.5463
26.3741 | -104.5418
26.3778 | -104.5361
26.3815 | -104.5306
26.3852 | -104.527
26.3889 | -104.5245
26.3926 | -104.5222
26.3963 | -104.5193
26.4 | -104.5171
26.4037 | -104.515
26.4074 | -104.5146
26.4111 | -104.5142
26.4148 | -104.5162
26.4185 | -104.5187
26.4222 | -104.5205
26.4259 | -104.5228
26.4296 | -104.5258
26.4333 | -104.5293
26.437 | -104.5328
26.4407 | -104.5363
26.4444 | -104.5419
26.4481 | -104.5478
26.4519 | -104.5556
26.4556 | -104.5653
26.4593 | -104.5743
26.463 | -104.5828
26.4667 | -104.5914
26.4704 | -104.5996
26.4741 | -104.6094
26.4778 | -104.6209
26.4815 | -104.6319
26.4852 | -104.6441
26.4889 | -104.6566
26.4926 | -104.6697
26.4963 | -104.6824
26.5 | -104.695
26.5037 | -104.7071
26.5074 | -104.721
26.5111 | -104.7352
26.5148 | -104.7492
26.5185 | -104.7631
26.5222 | -104.7775
26.5259 | -104.7917
26.5296 | -104.8053
26.5333 | -104.82
26.537 | -104.8334
26.5407 | -104.847
26.5444 | -104.86
26.5481 | -104.8727
26.5519 | -104.8872
26.5556 | -104.9008
26.5593 | -104.913
26.563 | -104.9244
26.5667 | -104.9362
26.5704 | -104.9462
26.5741 | -104.9558
26.5778 | -104.9662
26.5815 | -104.9753
26.5852 | -104.9821
26.5889 | -104.987
26.5926 | -104.9946
26.5963 | -105.0009
26.6 | -105.0062
26.6037 | -105.0095
26.6074 | -105.0135
26.6111 | -105.0172
26.6148 | -105.0189
26.6185 | -105.0176
26.6222 | -105.0165
26.6259 | -105.0139
26.6296 | -105.0117
26.6333 | -105.0096
26.637 | -105.0069
26.6407 | -105.0045
26.6444 | -105.0016
26.6481 | -104.996
26.6519 | -104.9909
26.6556 | -104.9867
26.6593 | -104.9808
26.663 | -104.9744
26.6667 | -104.9693
26.6704 | -104.9641
26.6741 | -104.9594
26.6778 | -104.9548
26.6815 | -104.9498
26.6852 | -104.9452
26.6889 | -104.94
26.6926 | -104.936
26.6963 | -104.9316
26.7 | -104.927
26.7037 | -104.923
26.7074 | -104.9177
26.7111 | -104.9133
26.7148 | -104.9103
26.7185 | -104.9063
26.7222 | -104.9017
26.7259 | -104.8986
26.7296 | -104.8942
26.7333 | -104.888
26.737 | -104.8829
26.7407 | -104.8798
26.7444 | -104.8778
26.7481 | -104.8759
26.7519 | -104.8725
26.7556 | -104.869
26.7593 | -104.8651
26.763 | -104.8611
26.7667 | -104.8576
26.7704 | -104.8558
26.7741 | -104.8538
26.7778 | -104.8519
26.7815 | -104.8494
26.7852 | -104.8467
26.7889 | -104.8441
26.7926 | -104.8386
26.7963 | -104.8342
26.8 | -104.8314
26.8037 | -104.8284
26.8074 | -104.8247
26.8111 | -104.8161
26.8148 | -104.8102
26.8185 | -104.8059
26.8222 | -104.8009
26.8259 | -104.7974
26.8296 | -104.7921
26.8333 | -104.7863
26.837 | -104.7812
26.8407 | -104.7755
26.8444 | -104.7704
26.8481 | -104.7668
26.8519 | -104.7627
26.8556 | -104.7586
26.8593 | -104.7539
26.863 | -104.748
26.8667 | -104.7404
26.8704 | -104.7337
26.8741 | -104.7286
26.8778 | -104.7242
26.8815 | -104.7208
26.8852 | -104.7147
26.8889 | -104.708
26.8926 | -104.7027
26.8963 | -104.6947
26.9 | -104.6865
26.9037 | -104.6785
26.9074 | -104.6706
26.9111 | -104.662
26.9148 | -104.6573
26.9185 | -104.6505
26.9222 | -104.6447
26.9259 | -104.639
26.9296 | -104.63
26.9333 | -104.6223
26.937 | -104.6146
26.9407 | -104.6077
26.9444 | -104.6023
26.9481 | -104.5961
26.9519 | -104.5895
26.9556 | -104.5831
26.9593 | -104.5776
26.963 | -104.5735
26.9667 | -104.5684
26.9704 | -104.561
26.9741 | -104.5525
26.9778 | -104.5457
26.9815 | -104.5396
26.9852 | -104.534
26.9889 | -104.5291
26.9926 | -104.5246
26.9963 | -104.5207
27 | -104.5172
27.0037 | -104.5143
27.0074 | -104.5106
27.0111 | -104.5069
27.0148 | -104.5042
27.0185 | -104.5013
27.0222 | -104.4978
27.0259 | -104.4956
27.0296 | -104.4928
27.0333 | -104.4884
27.037 | -104.4854
27.0407 | -104.4835
27.0444 | -104.4808
27.0481 | -104.4786
27.0519 | -104.4762
27.0556 | -104.474
27.0593 | -104.4723
27.063 | -104.471
27.0667 | -104.4692
27.0704 | -104.4674
27.0741 | -104.4654
27.0778 | -104.4629
27.0815 | -104.4607
27.0852 | -104.46
27.0889 | -104.4599
27.0926 | -104.4594
27.0963 | -104.4603
27.1 | -104.4623
27.1037 | -104.4632
27.1074 | -104.4635
27.1111 | -104.465
27.1148 | -104.467
27.1185 | -104.4689
27.1222 | -104.4714
27.1259 | -104.4725
27.1296 | -104.473
27.1333 | -104.4765
27.137 | -104.4812
27.1407 | -104.487
27.1444 | -104.4927
27.1481 | -104.4975
27.1519 | -104.5021
27.1556 | -104.5065
27.1593 | -104.5117
27.163 | -104.519
27.1667 | -104.5263
27.1704 | -104.5355
27.1741 | -104.5462
27.1778 | -104.557
27.1815 | -104.5666
27.1852 | -104.576
27.1889 | -104.5873
27.1926 | -104.602
27.1963 | -104.6133
27.2 | -104.623
27.2037 | -104.632
27.2074 | -104.6406
27.2111 | -104.6489
27.2148 | -104.658
27.2185 | -104.6658
27.2222 | -104.6746
27.2259 | -104.6837
27.2296 | -104.6919
27.2333 | -104.6995
27.237 | -104.7057
27.2407 | -104.7131
27.2444 | -104.7183
27.2481 | -104.7195
27.2519 | -104.7266
27.2556 | -104.7335
27.2593 | -104.7377
27.263 | -104.7393
27.2667 | -104.74
27.2704 | -104.7419
27.2741 | -104.7435
27.2778 | -104.7439
27.2815 | -104.7453
27.2852 | -104.7462
27.2889 | -104.7462
27.2926 | -104.7454
27.2963 | -104.7441
27.3 | -104.74
27.3037 | -104.7356
27.3074 | -104.7307
27.3111 | -104.7242
27.3148 | -104.7188
27.3185 | -104.712
27.3222 | -104.7067
27.3259 | -104.7011
27.3296 | -104.6931
27.3333 | -104.6858
27.337 | -104.6787
27.3407 | -104.6699
27.3444 | -104.6607
27.3481 | -104.6527
27.3519 | -104.6436
27.3556 | -104.6344
27.3593 | -104.6268
27.363 | -104.6176
27.3667 | -104.6067
27.3704 | -104.5965
27.3741 | -104.5879
27.3778 | -104.5787
27.3815 | -104.5695
27.3852 | -104.5595
27.3889 | -104.5488
27.3926 | -104.5384
27.3963 | -104.5297
27.4 | -104.5211
27.4037 | -104.5125
27.4074 | -104.505
27.4111 | -104.4973
27.4148 | -104.4887
27.4185 | -104.4806
27.4222 | -104.4722
27.4259 | -104.4644
27.4296 | -104.4564
27.4333 | -104.4492
27.437 | -104.4425
27.4407 | -104.4364
27.4444 | -104.4302
27.4481 | -104.4225
27.4519 | -104.4168
27.4556 | -104.4101
27.4593 | -104.4038
27.463 | -104.3991
27.4667 | -104.3941
27.4704 | -104.3898
27.4741 | -104.3843
27.4778 | -104.3783
27.4815 | -104.373
27.4852 | -104.3682
27.4889 | -104.3627
27.4926 | -104.3574
27.4963 | -104.3521
27.5 | -104.349
27.5037 | -104.345
27.5074 | -104.3406
27.5111 | -104.3381
27.5148 | -104.3371
27.5185 | -104.3347
27.5222 | -104.3326
27.5259 | -104.3293
27.5296 | -104.3263
27.5333 | -104.3264
27.537 | -104.3277
27.5407 | -104.3298
27.5444 | -104.3323
27.5481 | -104.3332
27.5519 | -104.334
27.5556 | -104.3351
27.5593 | -104.3377
27.563 | -104.3403
27.5667 | -104.3443
27.5704 | -104.3478
27.5741 | -104.3503
27.5778 | -104.3547
27.5815 | -104.3599
27.5852 | -104.3654
27.5889 | -104.3716
27.5926 | -104.3768
27.5963 | -104.382
27.6 | -104.3879
27.6037 | -104.3939
27.6074 | -104.4012
27.6111 | -104.408
27.6148 | -104.4143
27.6185 | -104.4207
27.6222 | -104.4291
27.6259 | -104.4367
27.6296 | -104.4458
27.6333 | -104.4527
27.637 | -104.4591
27.6407 | -104.4657
27.6444 | -104.471
27.6481 | -104.4762
27.6519 | -104.4825
27.6556 | -104.4885
27.6593 | -104.494
27.663 | -104.4994
27.6667 | -104.506
27.6704 | -104.5124
27.6741 | -104.5178
27.6778 | -104.5232
27.6815 | -104.5291
27.6852 | -104.5333
27.6889 | -104.5375
27.6926 | -104.542
27.6963 | -104.5457
27.7 | -104.5512
27.7037 | -104.5562
27.7074 | -104.56
27.7111 | -104.5622
27.7148 | -104.5642
27.7185 | -104.5665
27.7222 | -104.5677
27.7259 | -104.5674
27.7296 | -104.5681
27.7333 | -104.5694
27.737 | -104.5703
27.7407 | -104.5727
27.7444 | -104.5739
27.7481 | -104.5725
27.7519 | -104.571
27.7556 | -104.5717
27.7593 | -104.5715
27.763 | -104.5686
27.7667 | -104.5661
27.7704 | -104.5649
27.7741 | -104.5621
27.7778 | -104.5597
27.7815 | -104.5579
27.7852 | -104.5569
27.7889 | -104.5521
27.7926 | -104.5471
27.7963 | -104.5439
27.8 | -104.5412
27.8037 | -104.5342
27.8074 | -104.5303
27.8111 | -104.5276
27.8148 | -104.5239
27.8185 | -104.5201
27.8222 | -104.5172
27.8259 | -104.513
27.8296 | -104.5105
27.8333 | -104.5072
27.837 | -104.5034
27.8407 | -104.5014
27.8444 | -104.4995
27.8481 | -104.4977
27.8519 | -104.496
27.8556 | -104.4927
27.8593 | -104.4883
27.863 | -104.4832
27.8667 | -104.4795
27.8704 | -104.4775
27.8741 | -104.4757
27.8778 | -104.4743
27.8815 | -104.4737
27.8852 | -104.4729
27.8889 | -104.4724
27.8926 | -104.4707
27.8963 | -104.4684
27.9 | -104.4659
27.9037 | -104.4627
27.9074 | -104.4591
27.9111 | -104.4546
27.9148 | -104.4502
27.9185 | -104.4484
27.9222 | -104.4474
27.9259 | -104.444
27.9296 | -104.4398
27.9333 | -104.4375
27.937 | -104.4341
27.9407 | -104.4303
27.9444 | -104.4269
27.9481 | -104.4237
27.9519 | -104.4203
27.9556 | -104.4169
27.9593 | -104.4125
27.963 | -104.4094
27.9667 | -104.406
27.9704 | -104.4026
27.9741 | -104.3992
27.9778 | -104.396
27.9815 | -104.3924
27.9852 | -104.3897
27.9889 | -104.3878
27.9926 | -104.3847
27.9963 | -104.3807
28 | -104.3752
28.0037 | -104.3706