## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3d | 27 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 5
Hood | 02-91C/P | 16
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1001.1 | hPa |  | PRC PRO 01
Humidity | 31.6 | %/Hr
Internal temperature of calibration bench | 23.43 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1008 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 23.1 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.1 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.97 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3185 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Analog acquisition system | Analog Hood
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 06/03/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 15/03/2018
Technical Manager | Anthony HUE | Approval |  | 20/03/2018
CEO | Robert GLEMOT | Approval |  | 20/03/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.166 | 57.3261
0.012209 | -35.7403 | 38.0983
0.018313 | -34.6974 | 27.9639
0.024417 | -34.2513 | 21.6547
0.030521 | -34.1455 | 17.9437
0.036626 | -34.0105 | 14.8881
0.04273 | -33.9203 | 12.5056
0.048834 | -33.8758 | 10.9996
0.054938 | -33.8127 | 9.8657
0.061043 | -33.7477 | 8.7937
0.067147 | -33.7409 | 7.9363
0.073251 | -33.7457 | 7.1367
0.079355 | -33.7661 | 6.4552
0.08546 | -33.7811 | 5.8698
0.091564 | -33.7778 | 5.4428
0.097668 | -33.7611 | 5.1938
0.10377 | -33.741 | 4.933
0.10988 | -33.7253 | 4.6452
0.11598 | -33.7172 | 4.3422
0.12209 | -33.7201 | 4.0512
0.12819 | -33.7318 | 3.7688
0.13429 | -33.745 | 3.4904
0.1404 | -33.7477 | 3.2269
0.1465 | -33.7485 | 2.9807
0.15261 | -33.744 | 2.7679
0.15871 | -33.736 | 2.6078
0.16482 | -33.7277 | 2.4979
0.17092 | -33.7278 | 2.4441
0.17702 | -33.7275 | 2.3606
0.18313 | -33.7289 | 2.2736
0.18923 | -33.7308 | 2.1592
0.19534 | -33.7316 | 2.0296
0.20144 | -33.7318 | 1.8868
0.20754 | -33.7297 | 1.7628
0.21365 | -33.7297 | 1.6452
0.21975 | -33.7302 | 1.5435
0.22586 | -33.7323 | 1.4523
0.23196 | -33.7358 | 1.3766
0.23807 | -33.7409 | 1.3101
0.24417 | -33.7458 | 1.2466
0.25027 | -33.7506 | 1.1796
0.25638 | -33.7538 | 1.1117
0.26248 | -33.7563 | 1.0396
0.26859 | -33.7581 | 0.96856
0.27469 | -33.7594 | 0.8924
0.2808 | -33.7614 | 0.81798
0.2869 | -33.7633 | 0.74615
0.293 | -33.7652 | 0.67898
0.29911 | -33.767 | 0.62024
0.30521 | -33.7689 | 0.56273
0.31132 | -33.7702 | 0.52648
0.31742 | -33.7721 | 0.48599
0.32353 | -33.773 | 0.45146
0.32963 | -33.7742 | 0.42247
0.33573 | -33.7755 | 0.40025
0.34184 | -33.7772 | 0.37613
0.34794 | -33.7794 | 0.34631
0.35405 | -33.7818 | 0.317
0.36015 | -33.7843 | 0.28616
0.36626 | -33.7865 | 0.25249
0.37236 | -33.7888 | 0.21058
0.37846 | -33.7912 | 0.16214
0.38457 | -33.7923 | 0.10484
0.39067 | -33.7936 | 0.05604
0.39678 | -33.7944 | 0.0037485
0.40288 | -33.7951 | -0.049928
0.40899 | -33.7954 | -0.10118
0.41509 | -33.7952 | -0.15068
0.42119 | -33.7955 | -0.19364
0.4273 | -33.7959 | -0.23971
0.4334 | -33.7962 | -0.28593
0.43951 | -33.7969 | -0.33089
0.44561 | -33.798 | -0.37233
0.45172 | -33.7998 | -0.4084
0.45782 | -33.801 | -0.44828
0.46392 | -33.8025 | -0.48775
0.47003 | -33.8039 | -0.52407
0.47613 | -33.8051 | -0.5614
0.48224 | -33.806 | -0.59981
0.48834 | -33.8054 | -0.63819
0.49445 | -33.8056 | -0.67704
0.50055 | -33.8054 | -0.71294
0.50665 | -33.8051 | -0.75103
0.51276 | -33.8046 | -0.78868
0.51886 | -33.8039 | -0.82549
0.52497 | -33.8031 | -0.85354
0.53107 | -33.8027 | -0.88678
0.53717 | -33.8024 | -0.91843
0.54328 | -33.802 | -0.94533
0.54938 | -33.8022 | -0.97173
0.55549 | -33.8026 | -0.99624
0.56159 | -33.804 | -1.0117
0.5677 | -33.8046 | -1.0301
0.5738 | -33.806 | -1.05
0.5799 | -33.8072 | -1.0686
0.58601 | -33.8086 | -1.0824
0.59211 | -33.8104 | -1.0922
0.59822 | -33.812 | -1.105
0.60432 | -33.8138 | -1.1203
0.61043 | -33.8153 | -1.1316
0.61653 | -33.8168 | -1.1447
0.62263 | -33.8186 | -1.1571
0.62874 | -33.8195 | -1.17
0.63484 | -33.8206 | -1.1821
0.64095 | -33.8213 | -1.1948
0.64705 | -33.822 | -1.2069
0.65316 | -33.8224 | -1.2218
0.65926 | -33.8227 | -1.2403
0.66536 | -33.8218 | -1.2607
0.67147 | -33.8217 | -1.2767
0.67757 | -33.8216 | -1.2937
0.68368 | -33.8211 | -1.3151
0.68978 | -33.8207 | -1.3352
0.69589 | -33.8201 | -1.3567
0.70199 | -33.8196 | -1.3821
0.70809 | -33.8191 | -1.404
0.7142 | -33.8188 | -1.4269
0.7203 | -33.8181 | -1.449
0.72641 | -33.8178 | -1.4738
0.73251 | -33.8176 | -1.4966
0.73862 | -33.8182 | -1.5203
0.74472 | -33.8183 | -1.5428
0.75082 | -33.8185 | -1.567
0.75693 | -33.8192 | -1.5909
0.76303 | -33.8201 | -1.6145
0.76914 | -33.8222 | -1.636
0.77524 | -33.8235 | -1.6577
0.78135 | -33.825 | -1.6795
0.78745 | -33.8267 | -1.7009
0.79355 | -33.8282 | -1.7222
0.79966 | -33.8301 | -1.7431
0.80576 | -33.8325 | -1.7642
0.81187 | -33.8343 | -1.7872
0.81797 | -33.8363 | -1.8099
0.82408 | -33.8382 | -1.8318
0.83018 | -33.84 | -1.8539
0.83628 | -33.8421 | -1.8742
0.84239 | -33.8433 | -1.8969
0.84849 | -33.8449 | -1.9207
0.8546 | -33.8465 | -1.9425
0.8607 | -33.8478 | -1.965
0.86681 | -33.849 | -1.9863
0.87291 | -33.8501 | -2.0074
0.87901 | -33.8507 | -2.0285
0.88512 | -33.8516 | -2.0491
0.89122 | -33.8526 | -2.0701
0.89733 | -33.8532 | -2.0884
0.90343 | -33.8539 | -2.1087
0.90953 | -33.8544 | -2.1262
0.91564 | -33.8547 | -2.1448
0.92174 | -33.8553 | -2.1626
0.92785 | -33.8557 | -2.1801
0.93395 | -33.8562 | -2.1969
0.94006 | -33.8567 | -2.2125
0.94616 | -33.8569 | -2.2269
0.95226 | -33.8574 | -2.2425
0.95837 | -33.8579 | -2.2589
0.96447 | -33.8585 | -2.2744
0.97058 | -33.8588 | -2.2904
0.97668 | -33.8589 | -2.3065
0.98279 | -33.859 | -2.3216
0.98889 | -33.8592 | -2.3348
0.99499 | -33.8595 | -2.3508
1.0011 | -33.8597 | -2.3641
1.0072 | -33.8597 | -2.3781
1.0133 | -33.8597 | -2.3913
1.0194 | -33.8596 | -2.4066
1.0255 | -33.8596 | -2.4225
1.0316 | -33.8593 | -2.4367
1.0377 | -33.8591 | -2.4524
1.0438 | -33.8588 | -2.4675
1.0499 | -33.8582 | -2.4848
1.056 | -33.8577 | -2.5011
1.0621 | -33.8574 | -2.5175
1.0682 | -33.8571 | -2.5322
1.0743 | -33.8567 | -2.5496
1.0805 | -33.8563 | -2.5663
1.0866 | -33.8555 | -2.5834
1.0927 | -33.8548 | -2.5999
1.0988 | -33.8544 | -2.617
1.1049 | -33.8538 | -2.6334
1.111 | -33.8532 | -2.651
1.1171 | -33.8526 | -2.6674
1.1232 | -33.8523 | -2.6849
1.1293 | -33.852 | -2.7011
1.1354 | -33.8518 | -2.7184
1.1415 | -33.8516 | -2.7361
1.1476 | -33.8512 | -2.754
1.1537 | -33.8508 | -2.7709
1.1598 | -33.851 | -2.7882
1.1659 | -33.851 | -2.804
1.172 | -33.8509 | -2.8204
1.1781 | -33.8508 | -2.8369
1.1842 | -33.8508 | -2.8537
1.1903 | -33.8513 | -2.8675
1.1964 | -33.8514 | -2.8833
1.2025 | -33.8517 | -2.8998
1.2086 | -33.8521 | -2.915
1.2147 | -33.8523 | -2.9301
1.2209 | -33.8529 | -2.9441
1.227 | -33.8538 | -2.9575
1.2331 | -33.8545 | -2.9723
1.2392 | -33.855 | -2.987
1.2453 | -33.8557 | -3.0007
1.2514 | -33.8563 | -3.0135
1.2575 | -33.8572 | -3.0271
1.2636 | -33.8582 | -3.0389
1.2697 | -33.8591 | -3.0509
1.2758 | -33.86 | -3.0648
1.2819 | -33.8609 | -3.0775
1.288 | -33.8618 | -3.091
1.2941 | -33.8629 | -3.1022
1.3002 | -33.8638 | -3.113
1.3063 | -33.8646 | -3.1248
1.3124 | -33.8655 | -3.1364
1.3185 | -33.8664 | -3.1484
1.3246 | -33.8674 | -3.1603
1.3307 | -33.8684 | -3.1716
1.3368 | -33.8694 | -3.1828
1.3429 | -33.8703 | -3.1945
1.349 | -33.871 | -3.2065
1.3551 | -33.8719 | -3.2185
1.3613 | -33.8728 | -3.2295
1.3674 | -33.8735 | -3.2422
1.3735 | -33.8742 | -3.2545
1.3796 | -33.8749 | -3.2661
1.3857 | -33.8754 | -3.2778
1.3918 | -33.8762 | -3.2907
1.3979 | -33.8769 | -3.3024
1.404 | -33.8775 | -3.3164
1.4101 | -33.8782 | -3.3292
1.4162 | -33.8787 | -3.3422
1.4223 | -33.8794 | -3.3562
1.4284 | -33.88 | -3.3687
1.4345 | -33.8804 | -3.3826
1.4406 | -33.8808 | -3.3985
1.4467 | -33.8814 | -3.4126
1.4528 | -33.882 | -3.4265
1.4589 | -33.8827 | -3.4407
1.465 | -33.8831 | -3.454
1.4711 | -33.8837 | -3.4678
1.4772 | -33.8843 | -3.4835
1.4833 | -33.8849 | -3.4988
1.4894 | -33.8855 | -3.5135
1.4955 | -33.8862 | -3.5279
1.5016 | -33.8869 | -3.5428
1.5078 | -33.8874 | -3.5587
1.5139 | -33.8879 | -3.5741
1.52 | -33.8884 | -3.5889
1.5261 | -33.8889 | -3.6037
1.5322 | -33.8893 | -3.6202
1.5383 | -33.8898 | -3.6351
1.5444 | -33.8904 | -3.6499
1.5505 | -33.891 | -3.6655
1.5566 | -33.8917 | -3.6808
1.5627 | -33.8923 | -3.6968
1.5688 | -33.893 | -3.7127
1.5749 | -33.8935 | -3.7275
1.581 | -33.8939 | -3.7431
1.5871 | -33.8944 | -3.759
1.5932 | -33.8949 | -3.7741
1.5993 | -33.8955 | -3.7906
1.6054 | -33.8959 | -3.806
1.6115 | -33.8965 | -3.8223
1.6176 | -33.8969 | -3.8371
1.6237 | -33.8973 | -3.8531
1.6298 | -33.8977 | -3.8679
1.6359 | -33.8982 | -3.8828
1.642 | -33.8988 | -3.8971
1.6482 | -33.8992 | -3.9122
1.6543 | -33.8996 | -3.9281
1.6604 | -33.9001 | -3.9424
1.6665 | -33.9004 | -3.9571
1.6726 | -33.9008 | -3.9719
1.6787 | -33.9013 | -3.9866
1.6848 | -33.9014 | -4.0014
1.6909 | -33.9018 | -4.0163
1.697 | -33.9022 | -4.0307
1.7031 | -33.9024 | -4.0456
1.7092 | -33.9026 | -4.0599
1.7153 | -33.9029 | -4.0742
1.7214 | -33.9029 | -4.0886
1.7275 | -33.9031 | -4.1035
1.7336 | -33.9032 | -4.1186
1.7397 | -33.9036 | -4.1331
1.7458 | -33.9037 | -4.1474
1.7519 | -33.904 | -4.1616
1.758 | -33.9042 | -4.1757
1.7641 | -33.9045 | -4.1897
1.7702 | -33.9047 | -4.2047
1.7763 | -33.9048 | -4.2183
1.7824 | -33.905 | -4.2329
1.7885 | -33.9052 | -4.2476
1.7947 | -33.9056 | -4.2617
1.8008 | -33.9058 | -4.2751
1.8069 | -33.906 | -4.2889
1.813 | -33.9063 | -4.3033
1.8191 | -33.9065 | -4.3181
1.8252 | -33.9068 | -4.3328
1.8313 | -33.9071 | -4.3473
1.8374 | -33.9075 | -4.3613
1.8435 | -33.9078 | -4.3753
1.8496 | -33.908 | -4.3905
1.8557 | -33.9084 | -4.4055
1.8618 | -33.9089 | -4.4199
1.8679 | -33.9092 | -4.434
1.874 | -33.9096 | -4.4481
1.8801 | -33.9099 | -4.4628
1.8862 | -33.9104 | -4.4768
1.8923 | -33.9108 | -4.4911
1.8984 | -33.9111 | -4.5053
1.9045 | -33.9114 | -4.5191
1.9106 | -33.9117 | -4.5336
1.9167 | -33.9119 | -4.548
1.9228 | -33.9123 | -4.562
1.9289 | -33.9126 | -4.5763
1.9351 | -33.9129 | -4.5895
1.9412 | -33.9133 | -4.6029
1.9473 | -33.9138 | -4.6172
1.9534 | -33.9142 | -4.6309
1.9595 | -33.9145 | -4.6448
1.9656 | -33.915 | -4.6594
1.9717 | -33.9154 | -4.6734
1.9778 | -33.9159 | -4.6871
1.9839 | -33.9163 | -4.7009
1.99 | -33.9166 | -4.7143
1.9961 | -33.9169 | -4.7287
2.0022 | -33.9173 | -4.7425
2.0083 | -33.9176 | -4.7561
2.0144 | -33.918 | -4.7697
2.0205 | -33.9184 | -4.7837
2.0266 | -33.9189 | -4.7974
2.0327 | -33.9193 | -4.8113
2.0388 | -33.9198 | -4.8253
2.0449 | -33.9201 | -4.839
2.051 | -33.9205 | -4.8527
2.0571 | -33.921 | -4.8664
2.0632 | -33.9214 | -4.8802
2.0693 | -33.9217 | -4.8947
2.0754 | -33.9221 | -4.9083
2.0816 | -33.9225 | -4.9228
2.0877 | -33.9228 | -4.9367
2.0938 | -33.9233 | -4.9511
2.0999 | -33.9237 | -4.9651
2.106 | -33.9239 | -4.9791
2.1121 | -33.9243 | -4.9934
2.1182 | -33.9246 | -5.0066
2.1243 | -33.925 | -5.0198
2.1304 | -33.9253 | -5.0333
2.1365 | -33.9258 | -5.0473
2.1426 | -33.9261 | -5.0616
2.1487 | -33.9265 | -5.0752
2.1548 | -33.9271 | -5.0893
2.1609 | -33.9276 | -5.1032
2.167 | -33.9281 | -5.1167
2.1731 | -33.9285 | -5.1311
2.1792 | -33.9289 | -5.1445
2.1853 | -33.9294 | -5.1588
2.1914 | -33.9298 | -5.1724
2.1975 | -33.9303 | -5.1863
2.2036 | -33.9307 | -5.2008
2.2097 | -33.931 | -5.2146
2.2158 | -33.9315 | -5.2276
2.222 | -33.9319 | -5.2405
2.2281 | -33.9323 | -5.2543
2.2342 | -33.9327 | -5.2679
2.2403 | -33.933 | -5.281
2.2464 | -33.9335 | -5.294
2.2525 | -33.9338 | -5.3078
2.2586 | -33.9343 | -5.3207
2.2647 | -33.9348 | -5.3335
2.2708 | -33.9352 | -5.347
2.2769 | -33.9357 | -5.3601
2.283 | -33.936 | -5.3735
2.2891 | -33.9364 | -5.3868
2.2952 | -33.9367 | -5.3997
2.3013 | -33.9371 | -5.4129
2.3074 | -33.9375 | -5.4256
2.3135 | -33.9379 | -5.4394
2.3196 | -33.9383 | -5.4519
2.3257 | -33.9387 | -5.4649
2.3318 | -33.939 | -5.4783
2.3379 | -33.9393 | -5.4915
2.344 | -33.9396 | -5.5051
2.3501 | -33.94 | -5.5181
2.3562 | -33.9404 | -5.5308
2.3623 | -33.9407 | -5.5435
2.3685 | -33.941 | -5.5564
2.3746 | -33.9413 | -5.57
2.3807 | -33.9417 | -5.5827
2.3868 | -33.9421 | -5.5957
2.3929 | -33.9425 | -5.6087
2.399 | -33.9427 | -5.6215
2.4051 | -33.9432 | -5.6344
2.4112 | -33.9436 | -5.6472
2.4173 | -33.944 | -5.6601
2.4234 | -33.9442 | -5.6732
2.4295 | -33.9446 | -5.6862
2.4356 | -33.945 | -5.6994
2.4417 | -33.9453 | -5.7128
2.4478 | -33.9456 | -5.7259
2.4539 | -33.946 | -5.7385
2.46 | -33.9463 | -5.7521
2.4661 | -33.9467 | -5.7654
2.4722 | -33.947 | -5.7779
2.4783 | -33.9473 | -5.7906
2.4844 | -33.9476 | -5.8039
2.4905 | -33.9479 | -5.8171
2.4966 | -33.9483 | -5.83
2.5027 | -33.9486 | -5.8428
2.5089 | -33.9489 | -5.8558
2.515 | -33.9492 | -5.869
2.5211 | -33.9495 | -5.8819
2.5272 | -33.9498 | -5.895
2.5333 | -33.9501 | -5.9082
2.5394 | -33.9505 | -5.9212
2.5455 | -33.9508 | -5.9348
2.5516 | -33.9512 | -5.9473
2.5577 | -33.9516 | -5.9609
2.5638 | -33.9519 | -5.9748
2.5699 | -33.9522 | -5.9884
2.576 | -33.9524 | -6.0021
2.5821 | -33.9527 | -6.0156
2.5882 | -33.9531 | -6.0288
2.5943 | -33.9534 | -6.0419
2.6004 | -33.9538 | -6.0549
2.6065 | -33.9541 | -6.0675
2.6126 | -33.9544 | -6.0803
2.6187 | -33.9547 | -6.093
2.6248 | -33.955 | -6.1059
2.6309 | -33.9554 | -6.1194
2.637 | -33.9558 | -6.1332
2.6431 | -33.9561 | -6.1459
2.6492 | -33.9564 | -6.1593
2.6554 | -33.9568 | -6.1724
2.6615 | -33.9571 | -6.1856
2.6676 | -33.9575 | -6.1991
2.6737 | -33.9579 | -6.2124
2.6798 | -33.9584 | -6.2256
2.6859 | -33.9587 | -6.239
2.692 | -33.959 | -6.2521
2.6981 | -33.9594 | -6.2654
2.7042 | -33.9597 | -6.2787
2.7103 | -33.9601 | -6.2915
2.7164 | -33.9606 | -6.3046
2.7225 | -33.961 | -6.3182
2.7286 | -33.9614 | -6.3313
2.7347 | -33.9617 | -6.344
2.7408 | -33.9621 | -6.3563
2.7469 | -33.9625 | -6.3698
2.753 | -33.9629 | -6.3825
2.7591 | -33.9634 | -6.396
2.7652 | -33.9637 | -6.4096
2.7713 | -33.964 | -6.4225
2.7774 | -33.9644 | -6.4356
2.7835 | -33.9648 | -6.4488
2.7896 | -33.9651 | -6.4615
2.7958 | -33.9655 | -6.4744
2.8019 | -33.966 | -6.487
2.808 | -33.9664 | -6.5
2.8141 | -33.9668 | -6.5133
2.8202 | -33.9672 | -6.5263
2.8263 | -33.9677 | -6.5393
2.8324 | -33.9682 | -6.5528
2.8385 | -33.9687 | -6.5659
2.8446 | -33.9691 | -6.5787
2.8507 | -33.9696 | -6.5917
2.8568 | -33.9701 | -6.6044
2.8629 | -33.9705 | -6.6172
2.869 | -33.9709 | -6.6304
2.8751 | -33.9714 | -6.6433
2.8812 | -33.9719 | -6.6568
2.8873 | -33.9723 | -6.6698
2.8934 | -33.9728 | -6.683
2.8995 | -33.9733 | -6.6963
2.9056 | -33.9736 | -6.7096
2.9117 | -33.9741 | -6.7227
2.9178 | -33.9746 | -6.7356
2.9239 | -33.9751 | -6.7493
2.93 | -33.9757 | -6.763
2.9361 | -33.9761 | -6.7759
2.9423 | -33.9765 | -6.7888
2.9484 | -33.9771 | -6.802
2.9545 | -33.9776 | -6.8152
2.9606 | -33.9781 | -6.8288
2.9667 | -33.9786 | -6.8421
2.9728 | -33.979 | -6.8555
2.9789 | -33.9796 | -6.8687
2.985 | -33.9802 | -6.8822
2.9911 | -33.9806 | -6.8955
2.9972 | -33.9812 | -6.9089
3.0033 | -33.9818 | -6.9225
3.0094 | -33.9823 | -6.9358
3.0155 | -33.9829 | -6.9487
3.0216 | -33.9834 | -6.9621
3.0277 | -33.9838 | -6.976
3.0338 | -33.9844 | -6.9895
3.0399 | -33.9849 | -7.003
3.046 | -33.9853 | -7.0167
3.0521 | -33.9858 | -7.0299
3.0582 | -33.9862 | -7.0441
3.0643 | -33.9867 | -7.0577
3.0704 | -33.9872 | -7.0718
3.0765 | -33.9876 | -7.0851
3.0827 | -33.988 | -7.0984
3.0888 | -33.9885 | -7.1119
3.0949 | -33.989 | -7.1254
3.101 | -33.9895 | -7.1391
3.1071 | -33.99 | -7.1529
3.1132 | -33.9905 | -7.1668
3.1193 | -33.991 | -7.1804
3.1254 | -33.9915 | -7.194
3.1315 | -33.992 | -7.208
3.1376 | -33.9925 | -7.2216
3.1437 | -33.993 | -7.2352
3.1498 | -33.9934 | -7.2494
3.1559 | -33.9939 | -7.263
3.162 | -33.9943 | -7.2773
3.1681 | -33.9947 | -7.2914
3.1742 | -33.9952 | -7.305
3.1803 | -33.9957 | -7.3188
3.1864 | -33.9962 | -7.3323
3.1925 | -33.9967 | -7.3459
3.1986 | -33.9971 | -7.36
3.2047 | -33.9975 | -7.3743
3.2108 | -33.9979 | -7.3876
3.2169 | -33.9984 | -7.4015
3.223 | -33.9988 | -7.4154
3.2292 | -33.9991 | -7.4294
3.2353 | -33.9995 | -7.4424
3.2414 | -33.9999 | -7.4561
3.2475 | -34.0003 | -7.4696
3.2536 | -34.0007 | -7.4833
3.2597 | -34.0012 | -7.4965
3.2658 | -34.0015 | -7.5097
3.2719 | -34.0019 | -7.5229
3.278 | -34.0023 | -7.5367
3.2841 | -34.0027 | -7.55
3.2902 | -34.0031 | -7.5637
3.2963 | -34.0035 | -7.5773
3.3024 | -34.0038 | -7.5906
3.3085 | -34.0042 | -7.6044
3.3146 | -34.0045 | -7.6177
3.3207 | -34.005 | -7.6314
3.3268 | -34.0054 | -7.6451
3.3329 | -34.0057 | -7.6581
3.339 | -34.006 | -7.6711
3.3451 | -34.0064 | -7.6846
3.3512 | -34.0067 | -7.6975
3.3573 | -34.0071 | -7.711
3.3634 | -34.0075 | -7.7243
3.3696 | -34.0078 | -7.7379
3.3757 | -34.0081 | -7.7509
3.3818 | -34.0086 | -7.764
3.3879 | -34.0089 | -7.7773
3.394 | -34.0092 | -7.7909
3.4001 | -34.0096 | -7.8039
3.4062 | -34.0098 | -7.8171
3.4123 | -34.0101 | -7.8299
3.4184 | -34.0105 | -7.8433
3.4245 | -34.0109 | -7.8563
3.4306 | -34.0111 | -7.8693
3.4367 | -34.0115 | -7.8821
3.4428 | -34.0118 | -7.8947
3.4489 | -34.0121 | -7.9079
3.455 | -34.0124 | -7.9209
3.4611 | -34.0127 | -7.9343
3.4672 | -34.013 | -7.9476
3.4733 | -34.0134 | -7.9609
3.4794 | -34.0137 | -7.9738
3.4855 | -34.014 | -7.9865
3.4916 | -34.0144 | -7.9991
3.4977 | -34.0147 | -8.0126
3.5038 | -34.015 | -8.0257
3.5099 | -34.0153 | -8.0382
3.5161 | -34.0156 | -8.0507
3.5222 | -34.0159 | -8.0633
3.5283 | -34.0161 | -8.0764
3.5344 | -34.0164 | -8.0893
3.5405 | -34.0168 | -8.1021
3.5466 | -34.0171 | -8.1152
3.5527 | -34.0174 | -8.1281
3.5588 | -34.0176 | -8.1409
3.5649 | -34.0179 | -8.1537
3.571 | -34.0182 | -8.1665
3.5771 | -34.0185 | -8.1798
3.5832 | -34.0188 | -8.192
3.5893 | -34.0192 | -8.2047
3.5954 | -34.0195 | -8.2178
3.6015 | -34.0198 | -8.2308
3.6076 | -34.0202 | -8.2436
3.6137 | -34.0205 | -8.2559
3.6198 | -34.0209 | -8.2683
3.6259 | -34.0212 | -8.2813
3.632 | -34.0215 | -8.294
3.6381 | -34.0219 | -8.307
3.6442 | -34.0222 | -8.3197
3.6503 | -34.0225 | -8.3321
3.6565 | -34.0229 | -8.345
3.6626 | -34.0232 | -8.3576
3.6687 | -34.0236 | -8.3705
3.6748 | -34.024 | -8.3831
3.6809 | -34.0244 | -8.3959
3.687 | -34.0247 | -8.409
3.6931 | -34.0252 | -8.4213
3.6992 | -34.0255 | -8.4341
3.7053 | -34.0259 | -8.4467
3.7114 | -34.0262 | -8.4596
3.7175 | -34.0266 | -8.4724
3.7236 | -34.027 | -8.4845
3.7297 | -34.0274 | -8.4974
3.7358 | -34.0277 | -8.5104
3.7419 | -34.0281 | -8.5234
3.748 | -34.0285 | -8.5361
3.7541 | -34.0288 | -8.5489
3.7602 | -34.0293 | -8.5614
3.7663 | -34.0296 | -8.574
3.7724 | -34.03 | -8.5867
3.7785 | -34.0305 | -8.5992
3.7846 | -34.0308 | -8.6118
3.7907 | -34.0313 | -8.6248
3.7969 | -34.0317 | -8.6376
3.803 | -34.0321 | -8.6499
3.8091 | -34.0325 | -8.6626
3.8152 | -34.0329 | -8.6752
3.8213 | -34.0333 | -8.6881
3.8274 | -34.0338 | -8.7005
3.8335 | -34.0343 | -8.7133
3.8396 | -34.0347 | -8.726
3.8457 | -34.0352 | -8.7388
3.8518 | -34.0357 | -8.7515
3.8579 | -34.0362 | -8.7641
3.864 | -34.0366 | -8.7769
3.8701 | -34.0371 | -8.79
3.8762 | -34.0375 | -8.8025
3.8823 | -34.0379 | -8.8154
3.8884 | -34.0384 | -8.8281
3.8945 | -34.0389 | -8.8412
3.9006 | -34.0393 | -8.8542
3.9067 | -34.0398 | -8.8674
3.9128 | -34.0402 | -8.8802
3.9189 | -34.0407 | -8.8927
3.925 | -34.0411 | -8.9056
3.9311 | -34.0415 | -8.9179
3.9372 | -34.042 | -8.9308
3.9434 | -34.0425 | -8.9436
3.9495 | -34.043 | -8.9566
3.9556 | -34.0435 | -8.9692
3.9617 | -34.0439 | -8.9816
3.9678 | -34.0444 | -8.994
3.9739 | -34.0448 | -9.0068
3.98 | -34.0453 | -9.0197
3.9861 | -34.0457 | -9.0323
3.9922 | -34.0461 | -9.0451
3.9983 | -34.0466 | -9.0575
4.0044 | -34.0471 | -9.0698
4.0105 | -34.0475 | -9.0825
4.0166 | -34.048 | -9.0953
4.0227 | -34.0484 | -9.1081
4.0288 | -34.0488 | -9.1213
4.0349 | -34.0492 | -9.1337
4.041 | -34.0497 | -9.1464
4.0471 | -34.0501 | -9.1592
4.0532 | -34.0505 | -9.1717
4.0593 | -34.051 | -9.1846
4.0654 | -34.0515 | -9.1971
4.0715 | -34.0519 | -9.2102
4.0776 | -34.0524 | -9.2229
4.0838 | -34.0529 | -9.236
4.0899 | -34.0533 | -9.2485
4.096 | -34.0537 | -9.2611
4.1021 | -34.0542 | -9.2736
4.1082 | -34.0546 | -9.2861
4.1143 | -34.055 | -9.2985
4.1204 | -34.0555 | -9.311
4.1265 | -34.0559 | -9.3234
4.1326 | -34.0564 | -9.3357
4.1387 | -34.0569 | -9.3483
4.1448 | -34.0573 | -9.361
4.1509 | -34.0577 | -9.3736
4.157 | -34.0581 | -9.3866
4.1631 | -34.0585 | -9.3995
4.1692 | -34.0589 | -9.412
4.1753 | -34.0594 | -9.4245
4.1814 | -34.0598 | -9.4368
4.1875 | -34.0602 | -9.4497
4.1936 | -34.0607 | -9.4625
4.1997 | -34.0611 | -9.4748
4.2058 | -34.0615 | -9.4876
4.2119 | -34.0619 | -9.5001
4.218 | -34.0623 | -9.5124
4.2241 | -34.0627 | -9.5248
4.2303 | -34.0631 | -9.5373
4.2364 | -34.0635 | -9.5499
4.2425 | -34.0639 | -9.5625
4.2486 | -34.0643 | -9.5753
4.2547 | -34.0646 | -9.5876
4.2608 | -34.0649 | -9.6002
4.2669 | -34.0653 | -9.6125
4.273 | -34.0657 | -9.6252
4.2791 | -34.0661 | -9.6377
4.2852 | -34.0665 | -9.6504
4.2913 | -34.0668 | -9.6626
4.2974 | -34.0672 | -9.6755
4.3035 | -34.0676 | -9.6883
4.3096 | -34.068 | -9.701
4.3157 | -34.0684 | -9.7134
4.3218 | -34.0688 | -9.7258
4.3279 | -34.0692 | -9.7383
4.334 | -34.0695 | -9.7508
4.3401 | -34.0699 | -9.7629
4.3462 | -34.0703 | -9.7755
4.3523 | -34.0707 | -9.7876
4.3584 | -34.0711 | -9.7997
4.3645 | -34.0714 | -9.8121
4.3707 | -34.0718 | -9.8243
4.3768 | -34.0722 | -9.8367
4.3829 | -34.0726 | -9.8495
4.389 | -34.073 | -9.8619
4.3951 | -34.0733 | -9.8741
4.4012 | -34.0737 | -9.8869
4.4073 | -34.074 | -9.8991
4.4134 | -34.0744 | -9.9114
4.4195 | -34.0748 | -9.9238
4.4256 | -34.0753 | -9.9361
4.4317 | -34.0757 | -9.9484
4.4378 | -34.076 | -9.9605
4.4439 | -34.0764 | -9.9726
4.45 | -34.0767 | -9.985
4.4561 | -34.077 | -9.9973
4.4622 | -34.0773 | -10.0098
4.4683 | -34.0777 | -10.0221
4.4744 | -34.0781 | -10.0348
4.4805 | -34.0784 | -10.0468
4.4866 | -34.0788 | -10.0591
4.4927 | -34.0792 | -10.0717
4.4988 | -34.0796 | -10.0839
4.5049 | -34.08 | -10.0959
4.511 | -34.0804 | -10.108
4.5172 | -34.0807 | -10.1202
4.5233 | -34.0811 | -10.133
4.5294 | -34.0815 | -10.1451
4.5355 | -34.0819 | -10.1575
4.5416 | -34.0822 | -10.1697
4.5477 | -34.0826 | -10.1821
4.5538 | -34.0829 | -10.194
4.5599 | -34.0834 | -10.2063
4.566 | -34.0837 | -10.2185
4.5721 | -34.0841 | -10.2308
4.5782 | -34.0845 | -10.2434
4.5843 | -34.0849 | -10.2557
4.5904 | -34.0853 | -10.2681
4.5965 | -34.0857 | -10.2803
4.6026 | -34.0861 | -10.293
4.6087 | -34.0865 | -10.3052
4.6148 | -34.0869 | -10.3178
4.6209 | -34.0873 | -10.33
4.627 | -34.0877 | -10.3425
4.6331 | -34.088 | -10.3547
4.6392 | -34.0884 | -10.3671
4.6453 | -34.0888 | -10.3793
4.6514 | -34.0892 | -10.3917
4.6576 | -34.0895 | -10.4044
4.6637 | -34.0899 | -10.4163
4.6698 | -34.0902 | -10.4285
4.6759 | -34.0906 | -10.4412
4.682 | -34.091 | -10.4533
4.6881 | -34.0913 | -10.4656
4.6942 | -34.0917 | -10.4779
4.7003 | -34.0921 | -10.4903
4.7064 | -34.0925 | -10.5024
4.7125 | -34.093 | -10.5149
4.7186 | -34.0933 | -10.527
4.7247 | -34.0937 | -10.5394
4.7308 | -34.0941 | -10.5516
4.7369 | -34.0945 | -10.564
4.743 | -34.0949 | -10.5767
4.7491 | -34.0953 | -10.5893
4.7552 | -34.0957 | -10.6018
4.7613 | -34.096 | -10.6142
4.7674 | -34.0964 | -10.6264
4.7735 | -34.0968 | -10.6392
4.7796 | -34.0972 | -10.6517
4.7857 | -34.0976 | -10.6642
4.7918 | -34.098 | -10.6764
4.7979 | -34.0985 | -10.6887
4.8041 | -34.0988 | -10.7011
4.8102 | -34.0992 | -10.7131
4.8163 | -34.0996 | -10.7254
4.8224 | -34.1 | -10.7378
4.8285 | -34.1005 | -10.7498
4.8346 | -34.1009 | -10.762
4.8407 | -34.1012 | -10.7744
4.8468 | -34.1017 | -10.787
4.8529 | -34.1021 | -10.7995
4.859 | -34.1025 | -10.8122
4.8651 | -34.1029 | -10.8244
4.8712 | -34.1033 | -10.8365
4.8773 | -34.1037 | -10.8489
4.8834 | -34.1041 | -10.8613
4.8895 | -34.1045 | -10.8734
4.8956 | -34.1049 | -10.8858
4.9017 | -34.1053 | -10.898
4.9078 | -34.1057 | -10.9106
4.9139 | -34.1061 | -10.9226
4.92 | -34.1066 | -10.9346
4.9261 | -34.107 | -10.9467
4.9322 | -34.1074 | -10.9591
4.9383 | -34.1079 | -10.9714
4.9445 | -34.1083 | -10.9837
4.9506 | -34.1087 | -10.9958
4.9567 | -34.1091 | -11.0079
4.9628 | -34.1095 | -11.0202
4.9689 | -34.1099 | -11.0328
4.975 | -34.1103 | -11.0454
4.9811 | -34.1108 | -11.0575
4.9872 | -34.1112 | -11.0697
4.9933 | -34.1117 | -11.0822
4.9994 | -34.1121 | -11.0949
5.0055 | -34.1125 | -11.1073
5.0116 | -34.113 | -11.12
5.0177 | -34.1134 | -11.1326
5.0238 | -34.1139 | -11.1452
5.0299 | -34.1144 | -11.1579
5.036 | -34.1148 | -11.1702
5.0421 | -34.1152 | -11.1824
5.0482 | -34.1156 | -11.1948
5.0543 | -34.1161 | -11.207
5.0604 | -34.1165 | -11.2192
5.0665 | -34.117 | -11.2313
5.0726 | -34.1174 | -11.244
5.0787 | -34.1179 | -11.2566
5.0848 | -34.1183 | -11.2691
5.091 | -34.1187 | -11.2812
5.0971 | -34.1192 | -11.2932
5.1032 | -34.1197 | -11.3059
5.1093 | -34.1201 | -11.3182
5.1154 | -34.1205 | -11.3303
5.1215 | -34.1209 | -11.3425
5.1276 | -34.1214 | -11.3546
5.1337 | -34.1219 | -11.3666
5.1398 | -34.1223 | -11.3787
5.1459 | -34.1227 | -11.3913
5.152 | -34.1232 | -11.4038
5.1581 | -34.1236 | -11.4163
5.1642 | -34.124 | -11.4288
5.1703 | -34.1245 | -11.441
5.1764 | -34.125 | -11.4532
5.1825 | -34.1254 | -11.4653
5.1886 | -34.1258 | -11.4774
5.1947 | -34.1262 | -11.4897
5.2008 | -34.1267 | -11.5022
5.2069 | -34.1271 | -11.5147
5.213 | -34.1276 | -11.5268
5.2191 | -34.1281 | -11.5392
5.2252 | -34.1285 | -11.5516
5.2314 | -34.1289 | -11.5641
5.2375 | -34.1294 | -11.5761
5.2436 | -34.1298 | -11.5886
5.2497 | -34.1302 | -11.6007
5.2558 | -34.1307 | -11.6133
5.2619 | -34.1311 | -11.6259
5.268 | -34.1316 | -11.6384
5.2741 | -34.1321 | -11.6507
5.2802 | -34.1325 | -11.6635
5.2863 | -34.133 | -11.6757
5.2924 | -34.1334 | -11.6883
5.2985 | -34.1339 | -11.7007
5.3046 | -34.1344 | -11.7131
5.3107 | -34.1348 | -11.7257
5.3168 | -34.1353 | -11.7381
5.3229 | -34.1357 | -11.7505
5.329 | -34.1362 | -11.7629
5.3351 | -34.1366 | -11.775
5.3412 | -34.1371 | -11.7873
5.3473 | -34.1375 | -11.7997
5.3534 | -34.138 | -11.8122
5.3595 | -34.1384 | -11.8246
5.3656 | -34.1389 | -11.837
5.3717 | -34.1393 | -11.8497
5.3779 | -34.1398 | -11.8624
5.384 | -34.1402 | -11.8747
5.3901 | -34.1407 | -11.8874
5.3962 | -34.1411 | -11.8996
5.4023 | -34.1416 | -11.9122
5.4084 | -34.1421 | -11.9245
5.4145 | -34.1426 | -11.9366
5.4206 | -34.143 | -11.9491
5.4267 | -34.1435 | -11.9618
5.4328 | -34.1439 | -11.9743
5.4389 | -34.1444 | -11.9864
5.445 | -34.1448 | -11.9992
5.4511 | -34.1453 | -12.0113
5.4572 | -34.1458 | -12.0238
5.4633 | -34.1462 | -12.0362
5.4694 | -34.1467 | -12.0486
5.4755 | -34.1471 | -12.0608
5.4816 | -34.1476 | -12.0732
5.4877 | -34.1481 | -12.0854
5.4938 | -34.1485 | -12.0976
5.4999 | -34.149 | -12.1099
5.506 | -34.1495 | -12.1223
5.5121 | -34.1499 | -12.1346
5.5183 | -34.1504 | -12.1468
5.5244 | -34.1508 | -12.1594
5.5305 | -34.1513 | -12.1717
5.5366 | -34.1517 | -12.1839
5.5427 | -34.1522 | -12.1964
5.5488 | -34.1526 | -12.2088
5.5549 | -34.1531 | -12.2212
5.561 | -34.1535 | -12.2338
5.5671 | -34.154 | -12.246
5.5732 | -34.1545 | -12.2587
5.5793 | -34.155 | -12.2714
5.5854 | -34.1555 | -12.2837
5.5915 | -34.1559 | -12.2964
5.5976 | -34.1564 | -12.3089
5.6037 | -34.1568 | -12.3213
5.6098 | -34.1573 | -12.3337
5.6159 | -34.1577 | -12.3459
5.622 | -34.1582 | -12.3584
5.6281 | -34.1587 | -12.3709
5.6342 | -34.1591 | -12.3836
5.6403 | -34.1596 | -12.396
5.6464 | -34.1601 | -12.4085
5.6525 | -34.1605 | -12.421
5.6586 | -34.161 | -12.4332
5.6648 | -34.1615 | -12.4456
5.6709 | -34.162 | -12.4581
5.677 | -34.1625 | -12.4706
5.6831 | -34.1629 | -12.483
5.6892 | -34.1634 | -12.4957
5.6953 | -34.1639 | -12.508
5.7014 | -34.1643 | -12.5206
5.7075 | -34.1648 | -12.533
5.7136 | -34.1652 | -12.5453
5.7197 | -34.1657 | -12.5578
5.7258 | -34.1662 | -12.57
5.7319 | -34.1667 | -12.5824
5.738 | -34.1672 | -12.5948
5.7441 | -34.1677 | -12.6072
5.7502 | -34.1682 | -12.6193
5.7563 | -34.1686 | -12.6318
5.7624 | -34.1691 | -12.6444
5.7685 | -34.1696 | -12.6568
5.7746 | -34.17 | -12.6695
5.7807 | -34.1705 | -12.6819
5.7868 | -34.1709 | -12.6945
5.7929 | -34.1714 | -12.7069
5.799 | -34.1719 | -12.7192
5.8052 | -34.1723 | -12.7314
5.8113 | -34.1728 | -12.7437
5.8174 | -34.1733 | -12.756
5.8235 | -34.1738 | -12.7683
5.8296 | -34.1742 | -12.7809
5.8357 | -34.1747 | -12.7932
5.8418 | -34.1751 | -12.8058
5.8479 | -34.1756 | -12.8183
5.854 | -34.1761 | -12.8308
5.8601 | -34.1766 | -12.843
5.8662 | -34.177 | -12.8556
5.8723 | -34.1775 | -12.8678
5.8784 | -34.1779 | -12.8801
5.8845 | -34.1784 | -12.8924
5.8906 | -34.1789 | -12.9048
5.8967 | -34.1794 | -12.9171
5.9028 | -34.1799 | -12.9293
5.9089 | -34.1804 | -12.9416
5.915 | -34.1809 | -12.9537
5.9211 | -34.1813 | -12.9661
5.9272 | -34.1818 | -12.9782
5.9333 | -34.1823 | -12.9905
5.9394 | -34.1828 | -13.0028
5.9455 | -34.1833 | -13.0149
5.9517 | -34.1838 | -13.0271
5.9578 | -34.1843 | -13.0396
5.9639 | -34.1848 | -13.0516
5.97 | -34.1853 | -13.064
5.9761 | -34.1859 | -13.0762
5.9822 | -34.1863 | -13.0886
5.9883 | -34.1868 | -13.1009
5.9944 | -34.1873 | -13.1134
6.0005 | -34.1878 | -13.1258
6.0066 | -34.1883 | -13.138
6.0127 | -34.1888 | -13.1501
6.0188 | -34.1893 | -13.1626
6.0249 | -34.1898 | -13.175
6.031 | -34.1903 | -13.1872
6.0371 | -34.1908 | -13.1995
6.0432 | -34.1912 | -13.2117
6.0493 | -34.1917 | -13.2238
6.0554 | -34.1923 | -13.2359
6.0615 | -34.1928 | -13.2482
6.0676 | -34.1933 | -13.2607
6.0737 | -34.1938 | -13.2731
6.0798 | -34.1943 | -13.2853
6.0859 | -34.1948 | -13.2975
6.0921 | -34.1953 | -13.3096
6.0982 | -34.1958 | -13.3219
6.1043 | -34.1963 | -13.334
6.1104 | -34.1969 | -13.3465
6.1165 | -34.1974 | -13.3583
6.1226 | -34.1979 | -13.3705
6.1287 | -34.1984 | -13.3827
6.1348 | -34.1989 | -13.3949
6.1409 | -34.1994 | -13.4071
6.147 | -34.1999 | -13.4193
6.1531 | -34.2004 | -13.4315
6.1592 | -34.2009 | -13.4436
6.1653 | -34.2013 | -13.4557
6.1714 | -34.2018 | -13.4679
6.1775 | -34.2024 | -13.4803
6.1836 | -34.2028 | -13.4925
6.1897 | -34.2034 | -13.5047
6.1958 | -34.2039 | -13.5172
6.2019 | -34.2044 | -13.5295
6.208 | -34.2049 | -13.5417
6.2141 | -34.2054 | -13.5538
6.2202 | -34.206 | -13.5662
6.2263 | -34.2065 | -13.5783
6.2325 | -34.207 | -13.5901
6.2386 | -34.2075 | -13.602
6.2447 | -34.208 | -13.6145
6.2508 | -34.2085 | -13.6267
6.2569 | -34.2091 | -13.6388
6.263 | -34.2096 | -13.651
6.2691 | -34.2101 | -13.6631
6.2752 | -34.2105 | -13.6753
6.2813 | -34.211 | -13.6875
6.2874 | -34.2115 | -13.6995
6.2935 | -34.212 | -13.7116
6.2996 | -34.2125 | -13.7237
6.3057 | -34.213 | -13.7359
6.3118 | -34.2135 | -13.7482
6.3179 | -34.214 | -13.7603
6.324 | -34.2145 | -13.7724
6.3301 | -34.215 | -13.7845
6.3362 | -34.2156 | -13.7965
6.3423 | -34.2161 | -13.8088
6.3484 | -34.2166 | -13.821
6.3545 | -34.2171 | -13.8328
6.3606 | -34.2176 | -13.8447
6.3667 | -34.2181 | -13.8566
6.3728 | -34.2186 | -13.8688
6.379 | -34.2191 | -13.8809
6.3851 | -34.2196 | -13.8928
6.3912 | -34.2201 | -13.9049
6.3973 | -34.2206 | -13.9169
6.4034 | -34.2211 | -13.9288
6.4095 | -34.2216 | -13.9409
6.4156 | -34.2221 | -13.953
6.4217 | -34.2226 | -13.9651
6.4278 | -34.2231 | -13.9772
6.4339 | -34.2236 | -13.9892
6.44 | -34.2242 | -14.0012
6.4461 | -34.2247 | -14.0135
6.4522 | -34.2252 | -14.0253
6.4583 | -34.2256 | -14.0375
6.4644 | -34.2261 | -14.0497
6.4705 | -34.2266 | -14.0614
6.4766 | -34.2271 | -14.0735
6.4827 | -34.2276 | -14.0856
6.4888 | -34.2281 | -14.0974
6.4949 | -34.2286 | -14.1093
6.501 | -34.2291 | -14.1214
6.5071 | -34.2297 | -14.1331
6.5132 | -34.2302 | -14.1452
6.5194 | -34.2306 | -14.1573
6.5255 | -34.2312 | -14.1691
6.5316 | -34.2316 | -14.1812
6.5377 | -34.2321 | -14.1932
6.5438 | -34.2326 | -14.205
6.5499 | -34.2331 | -14.2168
6.556 | -34.2337 | -14.2288
6.5621 | -34.2341 | -14.2411
6.5682 | -34.2346 | -14.2531
6.5743 | -34.2352 | -14.265
6.5804 | -34.2357 | -14.2771
6.5865 | -34.2362 | -14.2891
6.5926 | -34.2367 | -14.301
6.5987 | -34.2372 | -14.3127
6.6048 | -34.2377 | -14.3247
6.6109 | -34.2383 | -14.3366
6.617 | -34.2387 | -14.3487
6.6231 | -34.2392 | -14.3606
6.6292 | -34.2397 | -14.3726
6.6353 | -34.2402 | -14.3848
6.6414 | -34.2408 | -14.3967
6.6475 | -34.2413 | -14.4086
6.6536 | -34.2418 | -14.4204
6.6597 | -34.2423 | -14.4321
6.6659 | -34.2428 | -14.444
6.672 | -34.2433 | -14.4558
6.6781 | -34.2438 | -14.4678
6.6842 | -34.2443 | -14.4795
6.6903 | -34.2449 | -14.4914
6.6964 | -34.2454 | -14.5034
6.7025 | -34.2459 | -14.5151
6.7086 | -34.2464 | -14.527
6.7147 | -34.2469 | -14.5389
6.7208 | -34.2474 | -14.5509
6.7269 | -34.2479 | -14.5628
6.733 | -34.2484 | -14.5745
6.7391 | -34.2489 | -14.5864
6.7452 | -34.2494 | -14.598
6.7513 | -34.2499 | -14.6096
6.7574 | -34.2504 | -14.6216
6.7635 | -34.251 | -14.6333
6.7696 | -34.2515 | -14.645
6.7757 | -34.252 | -14.6569
6.7818 | -34.2525 | -14.6688
6.7879 | -34.2531 | -14.6804
6.794 | -34.2536 | -14.6921
6.8001 | -34.2541 | -14.7039
6.8063 | -34.2546 | -14.7157
6.8124 | -34.2551 | -14.7277
6.8185 | -34.2556 | -14.7397
6.8246 | -34.256 | -14.7513
6.8307 | -34.2566 | -14.763
6.8368 | -34.2571 | -14.775
6.8429 | -34.2576 | -14.7869
6.849 | -34.2581 | -14.7987
6.8551 | -34.2587 | -14.8103
6.8612 | -34.2592 | -14.8221
6.8673 | -34.2597 | -14.8339
6.8734 | -34.2602 | -14.8458
6.8795 | -34.2607 | -14.8577
6.8856 | -34.2612 | -14.8695
6.8917 | -34.2617 | -14.8812
6.8978 | -34.2622 | -14.8927
6.9039 | -34.2627 | -14.9042
6.91 | -34.2633 | -14.9161
6.9161 | -34.2638 | -14.9276
6.9222 | -34.2643 | -14.9394
6.9283 | -34.2649 | -14.9508
6.9344 | -34.2654 | -14.9627
6.9405 | -34.2659 | -14.9744
6.9466 | -34.2664 | -14.9863
6.9528 | -34.267 | -14.9982
6.9589 | -34.2675 | -15.0101
6.965 | -34.268 | -15.0219
6.9711 | -34.2686 | -15.0337
6.9772 | -34.2691 | -15.0455
6.9833 | -34.2696 | -15.0572
6.9894 | -34.2701 | -15.0692
6.9955 | -34.2707 | -15.0806
7.0016 | -34.2712 | -15.0923
7.0077 | -34.2717 | -15.104
7.0138 | -34.2723 | -15.1156
7.0199 | -34.2727 | -15.1273
7.026 | -34.2733 | -15.139
7.0321 | -34.2738 | -15.1504
7.0382 | -34.2743 | -15.1618
7.0443 | -34.2749 | -15.1737
7.0504 | -34.2754 | -15.1855
7.0565 | -34.2759 | -15.1973
7.0626 | -34.2765 | -15.2087
7.0687 | -34.277 | -15.22
7.0748 | -34.2775 | -15.2314
7.0809 | -34.278 | -15.243
7.087 | -34.2786 | -15.2548
7.0932 | -34.2791 | -15.2664
7.0993 | -34.2796 | -15.2781
7.1054 | -34.2801 | -15.2896
7.1115 | -34.2807 | -15.3011
7.1176 | -34.2812 | -15.3127
7.1237 | -34.2817 | -15.3242
7.1298 | -34.2822 | -15.3358
7.1359 | -34.2827 | -15.3474
7.142 | -34.2833 | -15.3587
7.1481 | -34.2838 | -15.3703
7.1542 | -34.2843 | -15.3819
7.1603 | -34.2848 | -15.3935
7.1664 | -34.2854 | -15.4051
7.1725 | -34.2859 | -15.4168
7.1786 | -34.2865 | -15.4284
7.1847 | -34.287 | -15.44
7.1908 | -34.2875 | -15.4518
7.1969 | -34.288 | -15.4636
7.203 | -34.2885 | -15.475
7.2091 | -34.2891 | -15.4865
7.2152 | -34.2896 | -15.4981
7.2213 | -34.2901 | -15.5095
7.2274 | -34.2907 | -15.5211
7.2335 | -34.2912 | -15.5326
7.2397 | -34.2917 | -15.5441
7.2458 | -34.2923 | -15.5554
7.2519 | -34.2928 | -15.5671
7.258 | -34.2933 | -15.5788
7.2641 | -34.2939 | -15.5904
7.2702 | -34.2944 | -15.6021
7.2763 | -34.2949 | -15.6137
7.2824 | -34.2955 | -15.6254
7.2885 | -34.296 | -15.637
7.2946 | -34.2966 | -15.6486
7.3007 | -34.297 | -15.6603
7.3068 | -34.2976 | -15.672
7.3129 | -34.2981 | -15.6837
7.319 | -34.2987 | -15.6953
7.3251 | -34.2992 | -15.7067
7.3312 | -34.2997 | -15.7183
7.3373 | -34.3003 | -15.7297
7.3434 | -34.3008 | -15.7411
7.3495 | -34.3014 | -15.7526
7.3556 | -34.3019 | -15.7642
7.3617 | -34.3025 | -15.7759
7.3678 | -34.303 | -15.7874
7.3739 | -34.3035 | -15.7992
7.3801 | -34.304 | -15.811
7.3862 | -34.3046 | -15.8225
7.3923 | -34.3051 | -15.8342
7.3984 | -34.3057 | -15.8456
7.4045 | -34.3062 | -15.8571
7.4106 | -34.3068 | -15.8686
7.4167 | -34.3073 | -15.8802
7.4228 | -34.3079 | -15.8918
7.4289 | -34.3084 | -15.9034
7.435 | -34.309 | -15.9147
7.4411 | -34.3095 | -15.9263
7.4472 | -34.31 | -15.938
7.4533 | -34.3106 | -15.9494
7.4594 | -34.3111 | -15.9612
7.4655 | -34.3117 | -15.9725
7.4716 | -34.3122 | -15.9842
7.4777 | -34.3127 | -15.9958
7.4838 | -34.3133 | -16.0072
7.4899 | -34.3138 | -16.0187
7.496 | -34.3144 | -16.0302
7.5021 | -34.3149 | -16.0417
7.5082 | -34.3154 | -16.0532
7.5143 | -34.316 | -16.0648
7.5204 | -34.3165 | -16.0765
7.5266 | -34.3171 | -16.0879
7.5327 | -34.3176 | -16.0994
7.5388 | -34.3181 | -16.1112
7.5449 | -34.3187 | -16.1229
7.551 | -34.3192 | -16.1343
7.5571 | -34.3198 | -16.146
7.5632 | -34.3204 | -16.1576
7.5693 | -34.3209 | -16.169
7.5754 | -34.3215 | -16.1804
7.5815 | -34.322 | -16.192
7.5876 | -34.3225 | -16.2038
7.5937 | -34.3231 | -16.2154
7.5998 | -34.3237 | -16.2269
7.6059 | -34.3243 | -16.2387
7.612 | -34.3248 | -16.2502
7.6181 | -34.3254 | -16.2617
7.6242 | -34.3259 | -16.2735
7.6303 | -34.3265 | -16.2852
7.6364 | -34.327 | -16.2968
7.6425 | -34.3275 | -16.3084
7.6486 | -34.328 | -16.32
7.6547 | -34.3286 | -16.3315
7.6608 | -34.3292 | -16.343
7.667 | -34.3297 | -16.3547
7.6731 | -34.3303 | -16.3663
7.6792 | -34.3308 | -16.3781
7.6853 | -34.3314 | -16.3897
7.6914 | -34.3319 | -16.4014
7.6975 | -34.3325 | -16.4131
7.7036 | -34.333 | -16.4247
7.7097 | -34.3336 | -16.4365
7.7158 | -34.3342 | -16.4481
7.7219 | -34.3347 | -16.4596
7.728 | -34.3353 | -16.4713
7.7341 | -34.3358 | -16.483
7.7402 | -34.3364 | -16.4947
7.7463 | -34.3369 | -16.5063
7.7524 | -34.3375 | -16.5182
7.7585 | -34.338 | -16.5298
7.7646 | -34.3386 | -16.5414
7.7707 | -34.3391 | -16.5532
7.7768 | -34.3397 | -16.5649
7.7829 | -34.3402 | -16.5768
7.789 | -34.3408 | -16.5887
7.7951 | -34.3413 | -16.6003
7.8012 | -34.3419 | -16.6119
7.8073 | -34.3425 | -16.6237
7.8135 | -34.343 | -16.6354
7.8196 | -34.3436 | -16.6473
7.8257 | -34.3441 | -16.6591
7.8318 | -34.3447 | -16.6708
7.8379 | -34.3452 | -16.6824
7.844 | -34.3458 | -16.6939
7.8501 | -34.3463 | -16.7057
7.8562 | -34.3469 | -16.7176
7.8623 | -34.3475 | -16.7292
7.8684 | -34.348 | -16.741
7.8745 | -34.3486 | -16.7527
7.8806 | -34.3491 | -16.7645
7.8867 | -34.3497 | -16.7759
7.8928 | -34.3503 | -16.7874
7.8989 | -34.3508 | -16.7991
7.905 | -34.3514 | -16.8108
7.9111 | -34.3519 | -16.8225
7.9172 | -34.3525 | -16.8341
7.9233 | -34.353 | -16.8459
7.9294 | -34.3536 | -16.8576
7.9355 | -34.3542 | -16.8691
7.9416 | -34.3547 | -16.8808
7.9477 | -34.3552 | -16.8923
7.9539 | -34.3558 | -16.9039
7.96 | -34.3564 | -16.9158
7.9661 | -34.3569 | -16.9275
7.9722 | -34.3575 | -16.939
7.9783 | -34.3581 | -16.9509
7.9844 | -34.3587 | -16.9626
7.9905 | -34.3592 | -16.9745
7.9966 | -34.3598 | -16.9863
8.0027 | -34.3603 | -16.9979
8.0088 | -34.3609 | -17.0096
8.0149 | -34.3614 | -17.0213
8.021 | -34.362 | -17.0331
8.0271 | -34.3626 | -17.045
8.0332 | -34.3631 | -17.0566
8.0393 | -34.3637 | -17.0684
8.0454 | -34.3643 | -17.0804
8.0515 | -34.3649 | -17.0921
8.0576 | -34.3654 | -17.104
8.0637 | -34.366 | -17.1156
8.0698 | -34.3665 | -17.1271
8.0759 | -34.3671 | -17.1387
8.082 | -34.3676 | -17.1503
8.0881 | -34.3682 | -17.1621
8.0942 | -34.3688 | -17.1739
8.1004 | -34.3694 | -17.1857
8.1065 | -34.3699 | -17.1972
8.1126 | -34.3704 | -17.2089
8.1187 | -34.371 | -17.2208
8.1248 | -34.3716 | -17.2322
8.1309 | -34.3721 | -17.2438
8.137 | -34.3727 | -17.2553
8.1431 | -34.3733 | -17.2672
8.1492 | -34.3738 | -17.2789
8.1553 | -34.3744 | -17.2906
8.1614 | -34.375 | -17.3024
8.1675 | -34.3755 | -17.3141
8.1736 | -34.376 | -17.3259
8.1797 | -34.3766 | -17.3377
8.1858 | -34.3772 | -17.3495
8.1919 | -34.3777 | -17.3611
8.198 | -34.3783 | -17.3726
8.2041 | -34.3789 | -17.3844
8.2102 | -34.3795 | -17.396
8.2163 | -34.38 | -17.4076
8.2224 | -34.3806 | -17.4193
8.2285 | -34.3811 | -17.4307
8.2346 | -34.3817 | -17.4423
8.2408 | -34.3823 | -17.4538
8.2469 | -34.3828 | -17.4655
8.253 | -34.3834 | -17.4771
8.2591 | -34.384 | -17.4888
8.2652 | -34.3846 | -17.5005
8.2713 | -34.3852 | -17.5122
8.2774 | -34.3857 | -17.524
8.2835 | -34.3863 | -17.536
8.2896 | -34.3868 | -17.5474
8.2957 | -34.3874 | -17.5592
8.3018 | -34.388 | -17.5708
8.3079 | -34.3886 | -17.5823
8.314 | -34.3892 | -17.594
8.3201 | -34.3897 | -17.6055
8.3262 | -34.3903 | -17.6172
8.3323 | -34.3909 | -17.6288
8.3384 | -34.3914 | -17.6406
8.3445 | -34.392 | -17.6522
8.3506 | -34.3926 | -17.6639
8.3567 | -34.3932 | -17.6756
8.3628 | -34.3937 | -17.6873
8.3689 | -34.3943 | -17.6987
8.375 | -34.3949 | -17.7105
8.3812 | -34.3954 | -17.722
8.3873 | -34.396 | -17.7337
8.3934 | -34.3966 | -17.7454
8.3995 | -34.3971 | -17.757
8.4056 | -34.3977 | -17.7685
8.4117 | -34.3982 | -17.7803
8.4178 | -34.3988 | -17.792
8.4239 | -34.3994 | -17.8034
8.43 | -34.3999 | -17.8151
8.4361 | -34.4004 | -17.8269
8.4422 | -34.401 | -17.8383
8.4483 | -34.4016 | -17.8498
8.4544 | -34.4022 | -17.8615
8.4605 | -34.4028 | -17.873
8.4666 | -34.4033 | -17.8848
8.4727 | -34.4039 | -17.8965
8.4788 | -34.4045 | -17.908
8.4849 | -34.405 | -17.9196
8.491 | -34.4056 | -17.9312
8.4971 | -34.4062 | -17.9426
8.5032 | -34.4067 | -17.9541
8.5093 | -34.4073 | -17.9655
8.5154 | -34.4079 | -17.9771
8.5215 | -34.4085 | -17.9887
8.5277 | -34.409 | -18.0002
8.5338 | -34.4096 | -18.0118
8.5399 | -34.4101 | -18.0234
8.546 | -34.4107 | -18.0347
8.5521 | -34.4113 | -18.0463
8.5582 | -34.4119 | -18.0578
8.5643 | -34.4124 | -18.0691
8.5704 | -34.413 | -18.0804
8.5765 | -34.4136 | -18.0921
8.5826 | -34.4142 | -18.1033
8.5887 | -34.4148 | -18.1147
8.5948 | -34.4154 | -18.1263
8.6009 | -34.4159 | -18.1377
8.607 | -34.4165 | -18.1494
8.6131 | -34.4171 | -18.161
8.6192 | -34.4177 | -18.1722
8.6253 | -34.4182 | -18.1837
8.6314 | -34.4188 | -18.1951
8.6375 | -34.4193 | -18.2067
8.6436 | -34.4199 | -18.2181
8.6497 | -34.4205 | -18.2295
8.6558 | -34.421 | -18.2408
8.6619 | -34.4216 | -18.2523
8.6681 | -34.4222 | -18.2638
8.6742 | -34.4228 | -18.2753
8.6803 | -34.4233 | -18.2868
8.6864 | -34.4239 | -18.298
8.6925 | -34.4245 | -18.3094
8.6986 | -34.4251 | -18.3209
8.7047 | -34.4256 | -18.3325
8.7108 | -34.4262 | -18.344
8.7169 | -34.4268 | -18.3554
8.723 | -34.4273 | -18.3666
8.7291 | -34.4279 | -18.3779
8.7352 | -34.4285 | -18.3895
8.7413 | -34.429 | -18.401
8.7474 | -34.4296 | -18.4123
8.7535 | -34.4302 | -18.4238
8.7596 | -34.4308 | -18.4352
8.7657 | -34.4314 | -18.4467
8.7718 | -34.4319 | -18.4582
8.7779 | -34.4325 | -18.4697
8.784 | -34.4331 | -18.4811
8.7901 | -34.4337 | -18.4924
8.7962 | -34.4343 | -18.5036
8.8023 | -34.4348 | -18.515
8.8084 | -34.4354 | -18.5264
8.8146 | -34.436 | -18.5376
8.8207 | -34.4365 | -18.549
8.8268 | -34.4371 | -18.56
8.8329 | -34.4377 | -18.5714
8.839 | -34.4383 | -18.5828
8.8451 | -34.4389 | -18.594
8.8512 | -34.4395 | -18.6053
8.8573 | -34.44 | -18.6167
8.8634 | -34.4406 | -18.6278
8.8695 | -34.4412 | -18.6394
8.8756 | -34.4417 | -18.6507
8.8817 | -34.4423 | -18.662
8.8878 | -34.4429 | -18.6735
8.8939 | -34.4435 | -18.6849
8.9 | -34.4441 | -18.6961
8.9061 | -34.4446 | -18.7075
8.9122 | -34.4452 | -18.7188
8.9183 | -34.4458 | -18.7302
8.9244 | -34.4464 | -18.7417
8.9305 | -34.4471 | -18.7528
8.9366 | -34.4476 | -18.7641
8.9427 | -34.4482 | -18.7756
8.9488 | -34.4488 | -18.7869
8.955 | -34.4493 | -18.7983
8.9611 | -34.4499 | -18.8096
8.9672 | -34.4505 | -18.8206
8.9733 | -34.4511 | -18.8319
8.9794 | -34.4517 | -18.8432
8.9855 | -34.4523 | -18.8545
8.9916 | -34.4528 | -18.866
8.9977 | -34.4534 | -18.8773
9.0038 | -34.4541 | -18.8884
9.0099 | -34.4547 | -18.8998
9.016 | -34.4552 | -18.9112
9.0221 | -34.4558 | -18.9224
9.0282 | -34.4564 | -18.9336
9.0343 | -34.457 | -18.945
9.0404 | -34.4576 | -18.9562
9.0465 | -34.4582 | -18.9677
9.0526 | -34.4588 | -18.9789
9.0587 | -34.4594 | -18.9902
9.0648 | -34.46 | -19.0013
9.0709 | -34.4606 | -19.0126
9.077 | -34.4613 | -19.0239
9.0831 | -34.4618 | -19.0354
9.0892 | -34.4624 | -19.0467
9.0953 | -34.4631 | -19.0581
9.1015 | -34.4636 | -19.0693
9.1076 | -34.4643 | -19.0802
9.1137 | -34.4648 | -19.0915
9.1198 | -34.4655 | -19.1028
9.1259 | -34.4661 | -19.1141
9.132 | -34.4667 | -19.1254
9.1381 | -34.4673 | -19.1366
9.1442 | -34.4679 | -19.1476
9.1503 | -34.4685 | -19.1589
9.1564 | -34.4691 | -19.1701
9.1625 | -34.4697 | -19.1813
9.1686 | -34.4703 | -19.1924
9.1747 | -34.4709 | -19.2037
9.1808 | -34.4715 | -19.2148
9.1869 | -34.4721 | -19.2259
9.193 | -34.4727 | -19.2371
9.1991 | -34.4733 | -19.2484
9.2052 | -34.4739 | -19.2595
9.2113 | -34.4745 | -19.2707
9.2174 | -34.4752 | -19.2818
9.2235 | -34.4758 | -19.2929
9.2296 | -34.4764 | -19.304
9.2357 | -34.477 | -19.3153
9.2419 | -34.4776 | -19.3264
9.248 | -34.4783 | -19.3376
9.2541 | -34.4789 | -19.3489
9.2602 | -34.4795 | -19.3602
9.2663 | -34.4801 | -19.3711
9.2724 | -34.4807 | -19.3822
9.2785 | -34.4813 | -19.3933
9.2846 | -34.4819 | -19.4042
9.2907 | -34.4825 | -19.4153
9.2968 | -34.4831 | -19.4267
9.3029 | -34.4838 | -19.438
9.309 | -34.4844 | -19.4492
9.3151 | -34.485 | -19.4602
9.3212 | -34.4856 | -19.4712
9.3273 | -34.4862 | -19.4823
9.3334 | -34.4868 | -19.4933
9.3395 | -34.4874 | -19.5043
9.3456 | -34.488 | -19.5153
9.3517 | -34.4886 | -19.5266
9.3578 | -34.4892 | -19.5376
9.3639 | -34.4899 | -19.5486
9.37 | -34.4905 | -19.5597
9.3761 | -34.4911 | -19.5711
9.3822 | -34.4917 | -19.5822
9.3884 | -34.4924 | -19.5929
9.3945 | -34.493 | -19.6041
9.4006 | -34.4937 | -19.6152
9.4067 | -34.4943 | -19.6264
9.4128 | -34.4949 | -19.6377
9.4189 | -34.4955 | -19.649
9.425 | -34.4962 | -19.6599
9.4311 | -34.4968 | -19.671
9.4372 | -34.4974 | -19.6821
9.4433 | -34.498 | -19.6931
9.4494 | -34.4986 | -19.7041
9.4555 | -34.4993 | -19.7154
9.4616 | -34.4999 | -19.7266
9.4677 | -34.5005 | -19.7379
9.4738 | -34.5012 | -19.749
9.4799 | -34.5018 | -19.7601
9.486 | -34.5024 | -19.7712
9.4921 | -34.503 | -19.7823
9.4982 | -34.5037 | -19.7931
9.5043 | -34.5043 | -19.8043
9.5104 | -34.5049 | -19.8152
9.5165 | -34.5055 | -19.8263
9.5226 | -34.5062 | -19.8376
9.5288 | -34.5068 | -19.8483
9.5349 | -34.5074 | -19.8595
9.541 | -34.508 | -19.8704
9.5471 | -34.5086 | -19.8814
9.5532 | -34.5093 | -19.8926
9.5593 | -34.5099 | -19.9037
9.5654 | -34.5106 | -19.9145
9.5715 | -34.5112 | -19.9256
9.5776 | -34.5118 | -19.9369
9.5837 | -34.5124 | -19.948
9.5898 | -34.513 | -19.9592
9.5959 | -34.5136 | -19.9702
9.602 | -34.5142 | -19.9811
9.6081 | -34.5149 | -19.9921
9.6142 | -34.5155 | -20.0031
9.6203 | -34.5161 | -20.014
9.6264 | -34.5167 | -20.0251
9.6325 | -34.5173 | -20.0362
9.6386 | -34.518 | -20.0469
9.6447 | -34.5186 | -20.058
9.6508 | -34.5192 | -20.0689
9.6569 | -34.5199 | -20.0799
9.663 | -34.5205 | -20.091
9.6691 | -34.5212 | -20.1021
9.6753 | -34.5218 | -20.1132
9.6814 | -34.5224 | -20.124
9.6875 | -34.523 | -20.1349
9.6936 | -34.5236 | -20.1457
9.6997 | -34.5243 | -20.1568
9.7058 | -34.525 | -20.1674
9.7119 | -34.5256 | -20.1786
9.718 | -34.5262 | -20.1899
9.7241 | -34.5269 | -20.2007
9.7302 | -34.5274 | -20.2117
9.7363 | -34.5281 | -20.2226
9.7424 | -34.5287 | -20.2335
9.7485 | -34.5293 | -20.2445
9.7546 | -34.53 | -20.2554
9.7607 | -34.5306 | -20.2667
9.7668 | -34.5312 | -20.2777
9.7729 | -34.5318 | -20.2888
9.779 | -34.5325 | -20.2995
9.7851 | -34.5331 | -20.3106
9.7912 | -34.5338 | -20.3215
9.7973 | -34.5344 | -20.3326
9.8034 | -34.5351 | -20.3435
9.8095 | -34.5357 | -20.3546
9.8157 | -34.5364 | -20.3655
9.8218 | -34.537 | -20.3766
9.8279 | -34.5376 | -20.3875
9.834 | -34.5382 | -20.3986
9.8401 | -34.5389 | -20.4095
9.8462 | -34.5395 | -20.4202
9.8523 | -34.5401 | -20.4313
9.8584 | -34.5408 | -20.4423
9.8645 | -34.5414 | -20.4533
9.8706 | -34.542 | -20.4642
9.8767 | -34.5426 | -20.4751
9.8828 | -34.5432 | -20.486
9.8889 | -34.5438 | -20.4968
9.895 | -34.5445 | -20.5078
9.9011 | -34.5451 | -20.5187
9.9072 | -34.5457 | -20.5296
9.9133 | -34.5463 | -20.5406
9.9194 | -34.547 | -20.5514
9.9255 | -34.5476 | -20.5622
9.9316 | -34.5482 | -20.5731
9.9377 | -34.5489 | -20.584
9.9438 | -34.5495 | -20.595
9.9499 | -34.5501 | -20.6062
9.956 | -34.5507 | -20.6171
9.9622 | -34.5514 | -20.6282
9.9683 | -34.552 | -20.6391
9.9744 | -34.5526 | -20.6502
9.9805 | -34.5533 | -20.6612
9.9866 | -34.5539 | -20.6719
9.9927 | -34.5545 | -20.6827
9.9988 | -34.5551 | -20.6938
10.0049 | -34.5558 | -20.7046
10.011 | -34.5564 | -20.7157
10.0171 | -34.5571 | -20.7265
10.0232 | -34.5577 | -20.7373
10.0293 | -34.5584 | -20.7483
10.0354 | -34.559 | -20.7592
10.0415 | -34.5597 | -20.7702
10.0476 | -34.5603 | -20.7812
10.0537 | -34.5609 | -20.7922
10.0598 | -34.5615 | -20.8031
10.0659 | -34.5621 | -20.8141
10.072 | -34.5628 | -20.8248
10.0781 | -34.5634 | -20.8356
10.0842 | -34.564 | -20.8465
10.0903 | -34.5647 | -20.8574
10.0964 | -34.5653 | -20.868
10.1026 | -34.5659 | -20.8789
10.1087 | -34.5666 | -20.8899
10.1148 | -34.5672 | -20.9009
10.1209 | -34.5678 | -20.9116
10.127 | -34.5684 | -20.9222
10.1331 | -34.5691 | -20.933
10.1392 | -34.5697 | -20.9437
10.1453 | -34.5703 | -20.9546
10.1514 | -34.5709 | -20.9655
10.1575 | -34.5716 | -20.9766
10.1636 | -34.5722 | -20.9871
10.1697 | -34.5729 | -20.9981
10.1758 | -34.5735 | -21.0088
10.1819 | -34.5741 | -21.0197
10.188 | -34.5747 | -21.0306
10.1941 | -34.5753 | -21.0414
10.2002 | -34.5759 | -21.0522
10.2063 | -34.5766 | -21.063
10.2124 | -34.5772 | -21.0739
10.2185 | -34.5778 | -21.0849
10.2246 | -34.5785 | -21.0956
10.2307 | -34.5791 | -21.1064
10.2368 | -34.5797 | -21.1168
10.2429 | -34.5804 | -21.1277
10.2491 | -34.581 | -21.1387
10.2552 | -34.5816 | -21.1495
10.2613 | -34.5822 | -21.1604
10.2674 | -34.5829 | -21.1711
10.2735 | -34.5835 | -21.1821
10.2796 | -34.5842 | -21.1929
10.2857 | -34.5848 | -21.2038
10.2918 | -34.5854 | -21.2148
10.2979 | -34.586 | -21.2256
10.304 | -34.5867 | -21.2361
10.3101 | -34.5873 | -21.2471
10.3162 | -34.5879 | -21.258
10.3223 | -34.5886 | -21.2687
10.3284 | -34.5892 | -21.2794
10.3345 | -34.5898 | -21.2902
10.3406 | -34.5904 | -21.3008
10.3467 | -34.591 | -21.3115
10.3528 | -34.5917 | -21.3222
10.3589 | -34.5923 | -21.3331
10.365 | -34.5929 | -21.3439
10.3711 | -34.5935 | -21.3548
10.3772 | -34.5942 | -21.3655
10.3833 | -34.5948 | -21.3763
10.3895 | -34.5954 | -21.3872
10.3956 | -34.5961 | -21.398
10.4017 | -34.5967 | -21.4086
10.4078 | -34.5973 | -21.4194
10.4139 | -34.5979 | -21.43
10.42 | -34.5986 | -21.4409
10.4261 | -34.5992 | -21.4517
10.4322 | -34.5998 | -21.4625
10.4383 | -34.6004 | -21.4732
10.4444 | -34.6011 | -21.4837
10.4505 | -34.6017 | -21.4946
10.4566 | -34.6023 | -21.5052
10.4627 | -34.6029 | -21.5159
10.4688 | -34.6036 | -21.5266
10.4749 | -34.6042 | -21.5373
10.481 | -34.6048 | -21.548
10.4871 | -34.6054 | -21.5586
10.4932 | -34.6061 | -21.5693
10.4993 | -34.6067 | -21.5801
10.5054 | -34.6073 | -21.5909
10.5115 | -34.608 | -21.6017
10.5176 | -34.6086 | -21.6122
10.5237 | -34.6093 | -21.6231
10.5298 | -34.6099 | -21.6337
10.536 | -34.6105 | -21.6445
10.5421 | -34.6111 | -21.6552
10.5482 | -34.6118 | -21.666
10.5543 | -34.6124 | -21.6764
10.5604 | -34.613 | -21.687
10.5665 | -34.6137 | -21.6976
10.5726 | -34.6143 | -21.7083
10.5787 | -34.6149 | -21.719
10.5848 | -34.6155 | -21.7293
10.5909 | -34.6162 | -21.7401
10.597 | -34.6168 | -21.7509
10.6031 | -34.6174 | -21.7618
10.6092 | -34.618 | -21.7726
10.6153 | -34.6186 | -21.7834
10.6214 | -34.6193 | -21.7938
10.6275 | -34.6199 | -21.8046
10.6336 | -34.6205 | -21.8153
10.6397 | -34.6211 | -21.8261
10.6458 | -34.6218 | -21.8367
10.6519 | -34.6224 | -21.8475
10.658 | -34.623 | -21.8582
10.6641 | -34.6236 | -21.869
10.6702 | -34.6243 | -21.8797
10.6764 | -34.6249 | -21.8905
10.6825 | -34.6255 | -21.9011
10.6886 | -34.6262 | -21.9116
10.6947 | -34.6268 | -21.9222
10.7008 | -34.6274 | -21.9328
10.7069 | -34.6281 | -21.9436
10.713 | -34.6287 | -21.9543
10.7191 | -34.6293 | -21.9649
10.7252 | -34.6299 | -21.9752
10.7313 | -34.6306 | -21.9859
10.7374 | -34.6312 | -21.9966
10.7435 | -34.6318 | -22.0073
10.7496 | -34.6324 | -22.0178
10.7557 | -34.6331 | -22.0285
10.7618 | -34.6337 | -22.0389
10.7679 | -34.6343 | -22.0495
10.774 | -34.635 | -22.0603
10.7801 | -34.6356 | -22.071
10.7862 | -34.6362 | -22.0815
10.7923 | -34.6368 | -22.0922
10.7984 | -34.6375 | -22.1026
10.8045 | -34.6381 | -22.1131
10.8106 | -34.6387 | -22.1239
10.8168 | -34.6394 | -22.1345
10.8229 | -34.64 | -22.1452
10.829 | -34.6406 | -22.1557
10.8351 | -34.6413 | -22.166
10.8412 | -34.6419 | -22.1765
10.8473 | -34.6426 | -22.187
10.8534 | -34.6432 | -22.1976
10.8595 | -34.6438 | -22.2082
10.8656 | -34.6445 | -22.2186
10.8717 | -34.6451 | -22.2291
10.8778 | -34.6458 | -22.2398
10.8839 | -34.6464 | -22.2505
10.89 | -34.647 | -22.2612
10.8961 | -34.6477 | -22.2718
10.9022 | -34.6483 | -22.2822
10.9083 | -34.649 | -22.2927
10.9144 | -34.6496 | -22.3033
10.9205 | -34.6502 | -22.3138
10.9266 | -34.6509 | -22.3243
10.9327 | -34.6515 | -22.3348
10.9388 | -34.6522 | -22.3453
10.9449 | -34.6528 | -22.3558
10.951 | -34.6534 | -22.3663
10.9571 | -34.654 | -22.3769
10.9633 | -34.6547 | -22.3876
10.9694 | -34.6553 | -22.3983
10.9755 | -34.656 | -22.4088
10.9816 | -34.6566 | -22.4195
10.9877 | -34.6572 | -22.4298
10.9938 | -34.6579 | -22.4406
10.9999 | -34.6585 | -22.4509
11.006 | -34.6592 | -22.4612
11.0121 | -34.6598 | -22.4718
11.0182 | -34.6605 | -22.4823
11.0243 | -34.6611 | -22.4928
11.0304 | -34.6617 | -22.5034
11.0365 | -34.6624 | -22.5138
11.0426 | -34.6631 | -22.5242
11.0487 | -34.6637 | -22.5347
11.0548 | -34.6644 | -22.5451
11.0609 | -34.665 | -22.5558
11.067 | -34.6656 | -22.5663
11.0731 | -34.6663 | -22.5767
11.0792 | -34.667 | -22.5871
11.0853 | -34.6676 | -22.5976
11.0914 | -34.6682 | -22.6083
11.0975 | -34.6689 | -22.6188
11.1037 | -34.6695 | -22.6292
11.1098 | -34.6701 | -22.6398
11.1159 | -34.6708 | -22.6501
11.122 | -34.6714 | -22.6605
11.1281 | -34.6721 | -22.671
11.1342 | -34.6727 | -22.6815
11.1403 | -34.6734 | -22.6922
11.1464 | -34.674 | -22.7027
11.1525 | -34.6747 | -22.713
11.1586 | -34.6753 | -22.7237
11.1647 | -34.676 | -22.7341
11.1708 | -34.6766 | -22.7446
11.1769 | -34.6772 | -22.7551
11.183 | -34.6779 | -22.7653
11.1891 | -34.6786 | -22.7757
11.1952 | -34.6792 | -22.7861
11.2013 | -34.6799 | -22.7967
11.2074 | -34.6805 | -22.8071
11.2135 | -34.6811 | -22.8178
11.2196 | -34.6818 | -22.8282
11.2257 | -34.6825 | -22.8388
11.2318 | -34.6831 | -22.8492
11.2379 | -34.6838 | -22.8598
11.244 | -34.6845 | -22.8704
11.2502 | -34.6851 | -22.8808
11.2563 | -34.6858 | -22.8912
11.2624 | -34.6865 | -22.9016
11.2685 | -34.6871 | -22.912
11.2746 | -34.6878 | -22.9224
11.2807 | -34.6884 | -22.9329
11.2868 | -34.6891 | -22.9435
11.2929 | -34.6897 | -22.9537
11.299 | -34.6904 | -22.9641
11.3051 | -34.6911 | -22.9747
11.3112 | -34.6917 | -22.9853
11.3173 | -34.6924 | -22.9956
11.3234 | -34.6931 | -23.0058
11.3295 | -34.6937 | -23.0163
11.3356 | -34.6944 | -23.0268
11.3417 | -34.695 | -23.0371
11.3478 | -34.6956 | -23.0477
11.3539 | -34.6963 | -23.0582
11.36 | -34.6969 | -23.0684
11.3661 | -34.6976 | -23.0787
11.3722 | -34.6983 | -23.089
11.3783 | -34.6989 | -23.0995
11.3844 | -34.6996 | -23.11
11.3906 | -34.7003 | -23.1204
11.3967 | -34.7009 | -23.1307
11.4028 | -34.7016 | -23.1412
11.4089 | -34.7022 | -23.1515
11.415 | -34.7029 | -23.1619
11.4211 | -34.7035 | -23.1723
11.4272 | -34.7042 | -23.1828
11.4333 | -34.7048 | -23.193
11.4394 | -34.7055 | -23.2034
11.4455 | -34.7061 | -23.2137
11.4516 | -34.7068 | -23.2242
11.4577 | -34.7075 | -23.2346
11.4638 | -34.7082 | -23.2447
11.4699 | -34.7088 | -23.2552
11.476 | -34.7095 | -23.2658
11.4821 | -34.7102 | -23.2762
11.4882 | -34.7108 | -23.2867
11.4943 | -34.7115 | -23.2972
11.5004 | -34.7122 | -23.3073
11.5065 | -34.7129 | -23.3176
11.5126 | -34.7135 | -23.3279
11.5187 | -34.7142 | -23.3384
11.5248 | -34.7148 | -23.3489
11.5309 | -34.7155 | -23.3593
11.5371 | -34.7162 | -23.3695
11.5432 | -34.7169 | -23.38
11.5493 | -34.7175 | -23.3905
11.5554 | -34.7182 | -23.4009
11.5615 | -34.7189 | -23.4113
11.5676 | -34.7196 | -23.4217
11.5737 | -34.7202 | -23.432
11.5798 | -34.7209 | -23.4422
11.5859 | -34.7216 | -23.4525
11.592 | -34.7223 | -23.4628
11.5981 | -34.7229 | -23.4732
11.6042 | -34.7236 | -23.4834
11.6103 | -34.7243 | -23.4937
11.6164 | -34.7249 | -23.504
11.6225 | -34.7256 | -23.5143
11.6286 | -34.7263 | -23.5246
11.6347 | -34.7269 | -23.535
11.6408 | -34.7276 | -23.5452
11.6469 | -34.7283 | -23.5555
11.653 | -34.729 | -23.5658
11.6591 | -34.7296 | -23.5762
11.6652 | -34.7303 | -23.5865
11.6713 | -34.7309 | -23.5967
11.6775 | -34.7317 | -23.6069
11.6836 | -34.7323 | -23.6172
11.6897 | -34.733 | -23.6276
11.6958 | -34.7337 | -23.638
11.7019 | -34.7344 | -23.6485
11.708 | -34.7351 | -23.6589
11.7141 | -34.7358 | -23.669
11.7202 | -34.7365 | -23.6793
11.7263 | -34.7371 | -23.6897
11.7324 | -34.7378 | -23.7002
11.7385 | -34.7384 | -23.7106
11.7446 | -34.7391 | -23.7208
11.7507 | -34.7398 | -23.731
11.7568 | -34.7404 | -23.7413
11.7629 | -34.7411 | -23.7517
11.769 | -34.7418 | -23.7619
11.7751 | -34.7425 | -23.7721
11.7812 | -34.7432 | -23.7822
11.7873 | -34.7439 | -23.7925
11.7934 | -34.7445 | -23.8029
11.7995 | -34.7452 | -23.8132
11.8056 | -34.7459 | -23.8233
11.8117 | -34.7466 | -23.8336
11.8178 | -34.7473 | -23.8438
11.824 | -34.7479 | -23.854
11.8301 | -34.7486 | -23.8642
11.8362 | -34.7493 | -23.8744
11.8423 | -34.75 | -23.8847
11.8484 | -34.7506 | -23.895
11.8545 | -34.7514 | -23.905
11.8606 | -34.752 | -23.9152
11.8667 | -34.7527 | -23.9256
11.8728 | -34.7534 | -23.9357
11.8789 | -34.7541 | -23.9459
11.885 | -34.7547 | -23.9563
11.8911 | -34.7555 | -23.9664
11.8972 | -34.7562 | -23.9766
11.9033 | -34.7568 | -23.9868
11.9094 | -34.7575 | -23.997
11.9155 | -34.7582 | -24.0072
11.9216 | -34.7589 | -24.0172
11.9277 | -34.7596 | -24.0274
11.9338 | -34.7602 | -24.0377
11.9399 | -34.7609 | -24.048
11.946 | -34.7616 | -24.0582
11.9521 | -34.7623 | -24.0682
11.9582 | -34.763 | -24.0784
11.9644 | -34.7636 | -24.0884
11.9705 | -34.7644 | -24.0986
11.9766 | -34.7651 | -24.1088
11.9827 | -34.7658 | -24.119
11.9888 | -34.7665 | -24.1291
11.9949 | -34.7672 | -24.139
12.001 | -34.7679 | -24.1493
12.0071 | -34.7685 | -24.1595
12.0132 | -34.7692 | -24.1696
12.0193 | -34.7699 | -24.1798
12.0254 | -34.7706 | -24.19
12.0315 | -34.7713 | -24.1999
12.0376 | -34.7719 | -24.21
12.0437 | -34.7726 | -24.2201
12.0498 | -34.7733 | -24.2303
12.0559 | -34.774 | -24.2404
12.062 | -34.7747 | -24.2502
12.0681 | -34.7754 | -24.2605
12.0742 | -34.7761 | -24.2705
12.0803 | -34.7768 | -24.2809
12.0864 | -34.7775 | -24.2911
12.0925 | -34.7782 | -24.3012
12.0986 | -34.7789 | -24.3113
12.1047 | -34.7796 | -24.3215
12.1109 | -34.7803 | -24.3315
12.117 | -34.781 | -24.3416
12.1231 | -34.7817 | -24.3517
12.1292 | -34.7823 | -24.3618
12.1353 | -34.7831 | -24.3718
12.1414 | -34.7838 | -24.382
12.1475 | -34.7844 | -24.3922
12.1536 | -34.7851 | -24.4023
12.1597 | -34.7858 | -24.4124
12.1658 | -34.7865 | -24.4225
12.1719 | -34.7872 | -24.4324
12.178 | -34.7879 | -24.4426
12.1841 | -34.7886 | -24.4528
12.1902 | -34.7893 | -24.4629
12.1963 | -34.79 | -24.473
12.2024 | -34.7907 | -24.4828
12.2085 | -34.7914 | -24.493
12.2146 | -34.7921 | -24.5031
12.2207 | -34.7928 | -24.5132
12.2268 | -34.7935 | -24.5233
12.2329 | -34.7942 | -24.5332
12.239 | -34.795 | -24.543
12.2451 | -34.7957 | -24.5532
12.2513 | -34.7964 | -24.5633
12.2574 | -34.797 | -24.5735
12.2635 | -34.7978 | -24.5835
12.2696 | -34.7985 | -24.5935
12.2757 | -34.7992 | -24.6033
12.2818 | -34.7999 | -24.6134
12.2879 | -34.8006 | -24.6235
12.294 | -34.8013 | -24.6334
12.3001 | -34.802 | -24.6435
12.3062 | -34.8027 | -24.6537
12.3123 | -34.8034 | -24.6634
12.3184 | -34.8041 | -24.6733
12.3245 | -34.8048 | -24.6833
12.3306 | -34.8055 | -24.6933
12.3367 | -34.8062 | -24.7033
12.3428 | -34.8069 | -24.7133
12.3489 | -34.8076 | -24.7229
12.355 | -34.8083 | -24.7331
12.3611 | -34.809 | -24.7432
12.3672 | -34.8097 | -24.7532
12.3733 | -34.8104 | -24.7632
12.3794 | -34.8112 | -24.7729
12.3855 | -34.8118 | -24.7828
12.3916 | -34.8125 | -24.7929
12.3978 | -34.8133 | -24.803
12.4039 | -34.814 | -24.8131
12.41 | -34.8147 | -24.823
12.4161 | -34.8154 | -24.8327
12.4222 | -34.8161 | -24.8426
12.4283 | -34.8168 | -24.8525
12.4344 | -34.8175 | -24.8626
12.4405 | -34.8182 | -24.8726
12.4466 | -34.8189 | -24.8825
12.4527 | -34.8196 | -24.8923
12.4588 | -34.8203 | -24.9023
12.4649 | -34.821 | -24.9122
12.471 | -34.8217 | -24.9224
12.4771 | -34.8224 | -24.9324
12.4832 | -34.8231 | -24.9424
12.4893 | -34.8238 | -24.9521
12.4954 | -34.8245 | -24.962
12.5015 | -34.8252 | -24.9719
12.5076 | -34.826 | -24.9819
12.5137 | -34.8267 | -24.9918
12.5198 | -34.8274 | -25.0016
12.5259 | -34.8281 | -25.0116
12.532 | -34.8288 | -25.0217
12.5382 | -34.8295 | -25.0317
12.5443 | -34.8302 | -25.0416
12.5504 | -34.8309 | -25.0515
12.5565 | -34.8317 | -25.0612
12.5626 | -34.8324 | -25.0712
12.5687 | -34.8331 | -25.0812
12.5748 | -34.8338 | -25.0911
12.5809 | -34.8345 | -25.101
12.587 | -34.8352 | -25.111
12.5931 | -34.836 | -25.1207
12.5992 | -34.8367 | -25.1307
12.6053 | -34.8374 | -25.1407
12.6114 | -34.8381 | -25.1505
12.6175 | -34.8388 | -25.1609
12.6236 | -34.8395 | -25.1711
12.6297 | -34.8403 | -25.1805
12.6358 | -34.841 | -25.1903
12.6419 | -34.8417 | -25.2002
12.648 | -34.8424 | -25.2101
12.6541 | -34.8431 | -25.22
12.6602 | -34.8438 | -25.2295
12.6663 | -34.8445 | -25.2394
12.6724 | -34.8452 | -25.2491
12.6785 | -34.8459 | -25.2589
12.6847 | -34.8466 | -25.2687
12.6908 | -34.8473 | -25.2786
12.6969 | -34.8481 | -25.2882
12.703 | -34.8488 | -25.2982
12.7091 | -34.8495 | -25.308
12.7152 | -34.8502 | -25.3179
12.7213 | -34.8509 | -25.3278
12.7274 | -34.8516 | -25.3377
12.7335 | -34.8524 | -25.3472
12.7396 | -34.8531 | -25.3571
12.7457 | -34.8538 | -25.3672
12.7518 | -34.8545 | -25.3769
12.7579 | -34.8552 | -25.3868
12.764 | -34.8559 | -25.3967
12.7701 | -34.8566 | -25.4065
12.7762 | -34.8573 | -25.4163
12.7823 | -34.858 | -25.4261
12.7884 | -34.8588 | -25.4359
12.7945 | -34.8595 | -25.4459
12.8006 | -34.8602 | -25.4556
12.8067 | -34.8609 | -25.4653
12.8128 | -34.8616 | -25.475
12.8189 | -34.8623 | -25.4847
12.8251 | -34.863 | -25.4946
12.8312 | -34.8637 | -25.5045
12.8373 | -34.8644 | -25.514
12.8434 | -34.8652 | -25.5239
12.8495 | -34.8659 | -25.5337
12.8556 | -34.8665 | -25.5434
12.8617 | -34.8673 | -25.5533
12.8678 | -34.868 | -25.5632
12.8739 | -34.8687 | -25.5728
12.88 | -34.8694 | -25.5826
12.8861 | -34.8701 | -25.5922
12.8922 | -34.8708 | -25.602
12.8983 | -34.8715 | -25.6119
12.9044 | -34.8722 | -25.6215
12.9105 | -34.8729 | -25.631
12.9166 | -34.8736 | -25.6408
12.9227 | -34.8743 | -25.6505
12.9288 | -34.875 | -25.6604
12.9349 | -34.8758 | -25.6702
12.941 | -34.8765 | -25.6799
12.9471 | -34.8772 | -25.6897
12.9532 | -34.8779 | -25.6996
12.9593 | -34.8786 | -25.7095
12.9654 | -34.8793 | -25.7193
12.9716 | -34.8801 | -25.7292
12.9777 | -34.8808 | -25.7388
12.9838 | -34.8815 | -25.7486
12.9899 | -34.8822 | -25.7585
12.996 | -34.8829 | -25.7683
13.0021 | -34.8837 | -25.7779
13.0082 | -34.8844 | -25.7875
13.0143 | -34.8851 | -25.7969
13.0204 | -34.8858 | -25.8065
13.0265 | -34.8866 | -25.8164
13.0326 | -34.8873 | -25.8261
13.0387 | -34.888 | -25.836
13.0448 | -34.8887 | -25.8458
13.0509 | -34.8894 | -25.8554
13.057 | -34.8901 | -25.8652
13.0631 | -34.8909 | -25.8749
13.0692 | -34.8916 | -25.8846
13.0753 | -34.8923 | -25.8946
13.0814 | -34.893 | -25.9044
13.0875 | -34.8937 | -25.9139
13.0936 | -34.8944 | -25.9237
13.0997 | -34.8951 | -25.9334
13.1058 | -34.8958 | -25.9431
13.112 | -34.8966 | -25.9529
13.1181 | -34.8973 | -25.9625
13.1242 | -34.898 | -25.9722
13.1303 | -34.8987 | -25.9821
13.1364 | -34.8994 | -25.9918
13.1425 | -34.9001 | -26.0014
13.1486 | -34.9008 | -26.0111
13.1547 | -34.9015 | -26.0207
13.1608 | -34.9022 | -26.0304
13.1669 | -34.903 | -26.0399
13.173 | -34.9037 | -26.0496
13.1791 | -34.9044 | -26.0593
13.1852 | -34.9051 | -26.0688
13.1913 | -34.9058 | -26.0782
13.1974 | -34.9065 | -26.088
13.2035 | -34.9072 | -26.0976
13.2096 | -34.9079 | -26.1073
13.2157 | -34.9086 | -26.1168
13.2218 | -34.9094 | -26.1265
13.2279 | -34.9101 | -26.1358
13.234 | -34.9108 | -26.1455
13.2401 | -34.9115 | -26.1553
13.2462 | -34.9122 | -26.1649
13.2524 | -34.9129 | -26.1746
13.2585 | -34.9136 | -26.1841
13.2646 | -34.9143 | -26.1938
13.2707 | -34.9151 | -26.2035
13.2768 | -34.9158 | -26.2133
13.2829 | -34.9165 | -26.223
13.289 | -34.9172 | -26.2324
13.2951 | -34.9179 | -26.2419
13.3012 | -34.9187 | -26.2514
13.3073 | -34.9194 | -26.2611
13.3134 | -34.9201 | -26.2708
13.3195 | -34.9208 | -26.2804
13.3256 | -34.9215 | -26.2899
13.3317 | -34.9222 | -26.2993
13.3378 | -34.9229 | -26.309
13.3439 | -34.9236 | -26.3187
13.35 | -34.9243 | -26.3284
13.3561 | -34.925 | -26.338
13.3622 | -34.9257 | -26.3476
13.3683 | -34.9264 | -26.357
13.3744 | -34.9271 | -26.3665
13.3805 | -34.9278 | -26.3762
13.3866 | -34.9285 | -26.3856
13.3927 | -34.9292 | -26.3952
13.3989 | -34.9299 | -26.4046
13.405 | -34.9306 | -26.4141
13.4111 | -34.9313 | -26.4236
13.4172 | -34.932 | -26.4332
13.4233 | -34.9327 | -26.443
13.4294 | -34.9334 | -26.4527
13.4355 | -34.9341 | -26.4621
13.4416 | -34.9349 | -26.4715
13.4477 | -34.9356 | -26.4811
13.4538 | -34.9363 | -26.4906
13.4599 | -34.937 | -26.5001
13.466 | -34.9377 | -26.5097
13.4721 | -34.9384 | -26.5191
13.4782 | -34.9391 | -26.5287
13.4843 | -34.9398 | -26.5383
13.4904 | -34.9405 | -26.5478
13.4965 | -34.9412 | -26.5575
13.5026 | -34.9419 | -26.5672
13.5087 | -34.9426 | -26.5766
13.5148 | -34.9434 | -26.5863
13.5209 | -34.9441 | -26.5961
13.527 | -34.9448 | -26.6056
13.5331 | -34.9455 | -26.615
13.5393 | -34.9462 | -26.6243
13.5454 | -34.9469 | -26.6339
13.5515 | -34.9476 | -26.6436
13.5576 | -34.9483 | -26.6532
13.5637 | -34.949 | -26.6629
13.5698 | -34.9498 | -26.6724
13.5759 | -34.9505 | -26.6816
13.582 | -34.9512 | -26.6911
13.5881 | -34.9519 | -26.7006
13.5942 | -34.9526 | -26.7103
13.6003 | -34.9533 | -26.7198
13.6064 | -34.954 | -26.7295
13.6125 | -34.9547 | -26.7389
13.6186 | -34.9555 | -26.7484
13.6247 | -34.9562 | -26.758
13.6308 | -34.9569 | -26.7676
13.6369 | -34.9576 | -26.7772
13.643 | -34.9583 | -26.7868
13.6491 | -34.959 | -26.7961
13.6552 | -34.9597 | -26.8056
13.6613 | -34.9604 | -26.8151
13.6674 | -34.9611 | -26.8247
13.6735 | -34.9618 | -26.8343
13.6796 | -34.9625 | -26.8437
13.6858 | -34.9632 | -26.853
13.6919 | -34.964 | -26.8625
13.698 | -34.9647 | -26.8721
13.7041 | -34.9654 | -26.8818
13.7102 | -34.9661 | -26.8913
13.7163 | -34.9668 | -26.9006
13.7224 | -34.9675 | -26.9102
13.7285 | -34.9682 | -26.9197
13.7346 | -34.9689 | -26.9293
13.7407 | -34.9696 | -26.9387
13.7468 | -34.9703 | -26.9482
13.7529 | -34.971 | -26.9575
13.759 | -34.9717 | -26.967
13.7651 | -34.9724 | -26.9765
13.7712 | -34.9732 | -26.9861
13.7773 | -34.9739 | -26.9957
13.7834 | -34.9746 | -27.0052
13.7895 | -34.9753 | -27.0145
13.7956 | -34.976 | -27.0242
13.8017 | -34.9767 | -27.0337
13.8078 | -34.9774 | -27.0431
13.8139 | -34.9781 | -27.0524
13.82 | -34.9788 | -27.0618
13.8262 | -34.9795 | -27.0712
13.8323 | -34.9802 | -27.0807
13.8384 | -34.9809 | -27.0902
13.8445 | -34.9816 | -27.0996
13.8506 | -34.9823 | -27.1091
13.8567 | -34.9831 | -27.1185
13.8628 | -34.9837 | -27.128
13.8689 | -34.9845 | -27.1376
13.875 | -34.9852 | -27.1471
13.8811 | -34.9859 | -27.1565
13.8872 | -34.9865 | -27.1658
13.8933 | -34.9873 | -27.175
13.8994 | -34.988 | -27.1846
13.9055 | -34.9887 | -27.1941
13.9116 | -34.9894 | -27.2035
13.9177 | -34.9901 | -27.213
13.9238 | -34.9908 | -27.2225
13.9299 | -34.9915 | -27.2317
13.936 | -34.9922 | -27.2411
13.9421 | -34.9929 | -27.2506
13.9482 | -34.9935 | -27.26
13.9543 | -34.9942 | -27.2694
13.9604 | -34.9949 | -27.2788
13.9665 | -34.9957 | -27.2881
13.9727 | -34.9964 | -27.2976
13.9788 | -34.9971 | -27.307
13.9849 | -34.9978 | -27.3165
13.991 | -34.9985 | -27.3259
13.9971 | -34.9992 | -27.335
14.0032 | -34.9999 | -27.3444
14.0093 | -35.0006 | -27.3538
14.0154 | -35.0013 | -27.363
14.0215 | -35.002 | -27.3724
14.0276 | -35.0027 | -27.3819
14.0337 | -35.0034 | -27.3912
14.0398 | -35.0041 | -27.4005
14.0459 | -35.0049 | -27.4099
14.052 | -35.0056 | -27.4194
14.0581 | -35.0062 | -27.4289
14.0642 | -35.007 | -27.4383
14.0703 | -35.0077 | -27.4476
14.0764 | -35.0084 | -27.4568
14.0825 | -35.0091 | -27.4664
14.0886 | -35.0098 | -27.4758
14.0947 | -35.0105 | -27.4852
14.1008 | -35.0112 | -27.4945
14.1069 | -35.0119 | -27.5035
14.1131 | -35.0126 | -27.513
14.1192 | -35.0133 | -27.5223
14.1253 | -35.014 | -27.5317
14.1314 | -35.0147 | -27.5409
14.1375 | -35.0154 | -27.5502
14.1436 | -35.0162 | -27.5596
14.1497 | -35.0169 | -27.5689
14.1558 | -35.0175 | -27.5783
14.1619 | -35.0182 | -27.5876
14.168 | -35.0189 | -27.597
14.1741 | -35.0197 | -27.6062
14.1802 | -35.0204 | -27.6157
14.1863 | -35.0211 | -27.6249
14.1924 | -35.0218 | -27.6343
14.1985 | -35.0225 | -27.6434
14.2046 | -35.0232 | -27.6528
14.2107 | -35.0239 | -27.6618
14.2168 | -35.0246 | -27.6713
14.2229 | -35.0253 | -27.6806
14.229 | -35.0261 | -27.6901
14.2351 | -35.0268 | -27.6995
14.2412 | -35.0275 | -27.7088
14.2473 | -35.0282 | -27.7178
14.2534 | -35.0289 | -27.7269
14.2596 | -35.0295 | -27.7361
14.2657 | -35.0302 | -27.7454
14.2718 | -35.0309 | -27.7547
14.2779 | -35.0316 | -27.7641
14.284 | -35.0323 | -27.7734
14.2901 | -35.033 | -27.7826
14.2962 | -35.0338 | -27.7919
14.3023 | -35.0344 | -27.8014
14.3084 | -35.0351 | -27.8106
14.3145 | -35.0359 | -27.8199
14.3206 | -35.0366 | -27.8293
14.3267 | -35.0373 | -27.8386
14.3328 | -35.038 | -27.8479
14.3389 | -35.0387 | -27.8571
14.345 | -35.0394 | -27.8664
14.3511 | -35.0401 | -27.8757
14.3572 | -35.0408 | -27.8851
14.3633 | -35.0415 | -27.8945
14.3694 | -35.0422 | -27.9037
14.3755 | -35.043 | -27.9129
14.3816 | -35.0437 | -27.9224
14.3877 | -35.0444 | -27.9314
14.3938 | -35.0451 | -27.9409
14.4 | -35.0458 | -27.9501
14.4061 | -35.0465 | -27.9594
14.4122 | -35.0472 | -27.9688
14.4183 | -35.0479 | -27.978
14.4244 | -35.0486 | -27.9873
14.4305 | -35.0494 | -27.9965
14.4366 | -35.0501 | -28.0057
14.4427 | -35.0508 | -28.0152
14.4488 | -35.0515 | -28.0245
14.4549 | -35.0522 | -28.0335
14.461 | -35.053 | -28.0428
14.4671 | -35.0536 | -28.0521
14.4732 | -35.0544 | -28.0615
14.4793 | -35.0551 | -28.0708
14.4854 | -35.0558 | -28.0801
14.4915 | -35.0565 | -28.089
14.4976 | -35.0572 | -28.0981
14.5037 | -35.058 | -28.1073
14.5098 | -35.0587 | -28.1166
14.5159 | -35.0594 | -28.1259
14.522 | -35.0601 | -28.135
14.5281 | -35.0608 | -28.1439
14.5342 | -35.0615 | -28.1532
14.5403 | -35.0622 | -28.1624
14.5465 | -35.0629 | -28.1717
14.5526 | -35.0636 | -28.1811
14.5587 | -35.0643 | -28.1905
14.5648 | -35.0651 | -28.1994
14.5709 | -35.0658 | -28.2085
14.577 | -35.0665 | -28.2177
14.5831 | -35.0672 | -28.2268
14.5892 | -35.0679 | -28.236
14.5953 | -35.0686 | -28.245
14.6014 | -35.0693 | -28.2543
14.6075 | -35.07 | -28.2635
14.6136 | -35.0707 | -28.2727
14.6197 | -35.0714 | -28.282
14.6258 | -35.0721 | -28.2913
14.6319 | -35.0728 | -28.3002
14.638 | -35.0735 | -28.3094
14.6441 | -35.0742 | -28.3187
14.6502 | -35.0749 | -28.3281
14.6563 | -35.0757 | -28.3373
14.6624 | -35.0764 | -28.3465
14.6685 | -35.0771 | -28.3554
14.6746 | -35.0778 | -28.3646
14.6807 | -35.0785 | -28.3738
14.6869 | -35.0792 | -28.3829
14.693 | -35.0799 | -28.392
14.6991 | -35.0806 | -28.4012
14.7052 | -35.0814 | -28.4101
14.7113 | -35.0821 | -28.4194
14.7174 | -35.0828 | -28.4286
14.7235 | -35.0835 | -28.4376
14.7296 | -35.0842 | -28.4468
14.7357 | -35.085 | -28.4558
14.7418 | -35.0857 | -28.465
14.7479 | -35.0864 | -28.4741
14.754 | -35.0872 | -28.4833
14.7601 | -35.0879 | -28.4923
14.7662 | -35.0886 | -28.5014
14.7723 | -35.0893 | -28.5101
14.7784 | -35.09 | -28.5194
14.7845 | -35.0907 | -28.5285
14.7906 | -35.0915 | -28.5378
14.7967 | -35.0922 | -28.5469
14.8028 | -35.0929 | -28.556
14.8089 | -35.0936 | -28.5649
14.815 | -35.0944 | -28.574
14.8211 | -35.0951 | -28.5831
14.8272 | -35.0958 | -28.5923
14.8334 | -35.0965 | -28.6015
14.8395 | -35.0972 | -28.6105
14.8456 | -35.098 | -28.6194
14.8517 | -35.0987 | -28.6286
14.8578 | -35.0994 | -28.6378
14.8639 | -35.1001 | -28.647
14.87 | -35.1008 | -28.6562
14.8761 | -35.1016 | -28.6651
14.8822 | -35.1023 | -28.6743
14.8883 | -35.103 | -28.6834
14.8944 | -35.1037 | -28.6924
14.9005 | -35.1044 | -28.7015
14.9066 | -35.1051 | -28.7107
14.9127 | -35.1059 | -28.7196
14.9188 | -35.1066 | -28.7286
14.9249 | -35.1073 | -28.7377
14.931 | -35.108 | -28.7467
14.9371 | -35.1087 | -28.7557
14.9432 | -35.1095 | -28.7646
14.9493 | -35.1102 | -28.7736
14.9554 | -35.1109 | -28.7825
14.9615 | -35.1116 | -28.7916
14.9676 | -35.1123 | -28.801
14.9738 | -35.113 | -28.8101
14.9799 | -35.1137 | -28.8194
14.986 | -35.1144 | -28.8283
14.9921 | -35.1151 | -28.8374
14.9982 | -35.1158 | -28.8465
15.0043 | -35.1165 | -28.8554
15.0104 | -35.1172 | -28.8647
15.0165 | -35.118 | -28.874
15.0226 | -35.1187 | -28.8828
15.0287 | -35.1194 | -28.8921
15.0348 | -35.1201 | -28.901
15.0409 | -35.1208 | -28.9101
15.047 | -35.1215 | -28.9193
15.0531 | -35.1223 | -28.9281
15.0592 | -35.123 | -28.9371
15.0653 | -35.1237 | -28.9461
15.0714 | -35.1245 | -28.9552
15.0775 | -35.1252 | -28.9641
15.0836 | -35.1259 | -28.9732
15.0897 | -35.1267 | -28.9819
15.0958 | -35.1274 | -28.9909
15.1019 | -35.1282 | -28.9999
15.108 | -35.1289 | -29.0088
15.1141 | -35.1296 | -29.0177
15.1203 | -35.1303 | -29.0268
15.1264 | -35.1311 | -29.0357
15.1325 | -35.1318 | -29.0447
15.1386 | -35.1325 | -29.0536
15.1447 | -35.1333 | -29.0626
15.1508 | -35.134 | -29.0716
15.1569 | -35.1347 | -29.0807
15.163 | -35.1355 | -29.0894
15.1691 | -35.1362 | -29.0984
15.1752 | -35.1369 | -29.1073
15.1813 | -35.1376 | -29.1163
15.1874 | -35.1384 | -29.1253
15.1935 | -35.1391 | -29.1341
15.1996 | -35.1398 | -29.1431
15.2057 | -35.1406 | -29.1521
15.2118 | -35.1413 | -29.1612
15.2179 | -35.142 | -29.1701
15.224 | -35.1427 | -29.179
15.2301 | -35.1435 | -29.1876
15.2362 | -35.1442 | -29.1966
15.2423 | -35.145 | -29.2057
15.2484 | -35.1457 | -29.2147
15.2545 | -35.1464 | -29.2238
15.2607 | -35.1471 | -29.2328
15.2668 | -35.1479 | -29.2414
15.2729 | -35.1486 | -29.2503
15.279 | -35.1493 | -29.2594
15.2851 | -35.1501 | -29.2685
15.2912 | -35.1508 | -29.2774
15.2973 | -35.1515 | -29.2865
15.3034 | -35.1522 | -29.2953
15.3095 | -35.153 | -29.3042
15.3156 | -35.1537 | -29.313
15.3217 | -35.1545 | -29.3219
15.3278 | -35.1552 | -29.331
15.3339 | -35.1559 | -29.3396
15.34 | -35.1567 | -29.3485
15.3461 | -35.1574 | -29.3574
15.3522 | -35.1581 | -29.3662
15.3583 | -35.1588 | -29.3753
15.3644 | -35.1596 | -29.384
15.3705 | -35.1603 | -29.3925
15.3766 | -35.1611 | -29.4014
15.3827 | -35.1618 | -29.4103
15.3888 | -35.1625 | -29.4192
15.3949 | -35.1633 | -29.4282
15.401 | -35.164 | -29.4372
15.4072 | -35.1647 | -29.446
15.4133 | -35.1654 | -29.4549
15.4194 | -35.1662 | -29.4638
15.4255 | -35.1669 | -29.4727
15.4316 | -35.1676 | -29.4817
15.4377 | -35.1684 | -29.4908
15.4438 | -35.1691 | -29.4996
15.4499 | -35.1698 | -29.5084
15.456 | -35.1706 | -29.5172
15.4621 | -35.1713 | -29.526
15.4682 | -35.172 | -29.535
15.4743 | -35.1728 | -29.5436
15.4804 | -35.1735 | -29.5525
15.4865 | -35.1742 | -29.5614
15.4926 | -35.1749 | -29.5703
15.4987 | -35.1757 | -29.5793
15.5048 | -35.1764 | -29.5883
15.5109 | -35.1772 | -29.597
15.517 | -35.1779 | -29.6059
15.5231 | -35.1787 | -29.6148
15.5292 | -35.1794 | -29.6236
15.5353 | -35.1801 | -29.6326
15.5414 | -35.1808 | -29.6414
15.5476 | -35.1816 | -29.6499
15.5537 | -35.1824 | -29.6587
15.5598 | -35.1831 | -29.6677
15.5659 | -35.1838 | -29.6766
15.572 | -35.1846 | -29.6854
15.5781 | -35.1853 | -29.694
15.5842 | -35.1861 | -29.7027
15.5903 | -35.1868 | -29.7114
15.5964 | -35.1875 | -29.7202
15.6025 | -35.1883 | -29.729
15.6086 | -35.189 | -29.738
15.6147 | -35.1898 | -29.747
15.6208 | -35.1905 | -29.7554
15.6269 | -35.1913 | -29.7642
15.633 | -35.192 | -29.7731
15.6391 | -35.1928 | -29.7822
15.6452 | -35.1935 | -29.791
15.6513 | -35.1942 | -29.7995
15.6574 | -35.195 | -29.8084
15.6635 | -35.1957 | -29.8173
15.6696 | -35.1964 | -29.8261
15.6757 | -35.1972 | -29.835
15.6818 | -35.1979 | -29.8437
15.688 | -35.1986 | -29.8525
15.6941 | -35.1994 | -29.8613
15.7002 | -35.2001 | -29.8701
15.7063 | -35.2008 | -29.879
15.7124 | -35.2016 | -29.8879
15.7185 | -35.2023 | -29.8966
15.7246 | -35.2031 | -29.9052
15.7307 | -35.2038 | -29.9139
15.7368 | -35.2046 | -29.9229
15.7429 | -35.2053 | -29.9317
15.749 | -35.206 | -29.9406
15.7551 | -35.2067 | -29.9494
15.7612 | -35.2075 | -29.9586
15.7673 | -35.2082 | -29.9672
15.7734 | -35.2089 | -29.9764
15.7795 | -35.2096 | -29.9854
15.7856 | -35.2104 | -29.9942
15.7917 | -35.2111 | -30.0028
15.7978 | -35.2118 | -30.012
15.8039 | -35.2126 | -30.0208
15.81 | -35.2133 | -30.0296
15.8161 | -35.2141 | -30.0383
15.8222 | -35.2148 | -30.0467
15.8283 | -35.2156 | -30.0551
15.8345 | -35.2163 | -30.0638
15.8406 | -35.217 | -30.0724
15.8467 | -35.2178 | -30.0814
15.8528 | -35.2186 | -30.0902
15.8589 | -35.2193 | -30.0986
15.865 | -35.2201 | -30.1072
15.8711 | -35.2208 | -30.116
15.8772 | -35.2215 | -30.1246
15.8833 | -35.2223 | -30.1335
15.8894 | -35.223 | -30.1424
15.8955 | -35.2237 | -30.1511
15.9016 | -35.2244 | -30.1594
15.9077 | -35.2252 | -30.1683
15.9138 | -35.2259 | -30.1771
15.9199 | -35.2266 | -30.1858
15.926 | -35.2273 | -30.1946
15.9321 | -35.2281 | -30.2033
15.9382 | -35.2288 | -30.2119
15.9443 | -35.2296 | -30.2206
15.9504 | -35.2303 | -30.2293
15.9565 | -35.231 | -30.2381
15.9626 | -35.2318 | -30.2467
15.9687 | -35.2325 | -30.2551
15.9749 | -35.2332 | -30.2636
15.981 | -35.234 | -30.2721
15.9871 | -35.2347 | -30.2808
15.9932 | -35.2355 | -30.2894
15.9993 | -35.2362 | -30.2983
16.0054 | -35.237 | -30.3067
16.0115 | -35.2377 | -30.3153
16.0176 | -35.2384 | -30.3242
16.0237 | -35.2392 | -30.3329
16.0298 | -35.2399 | -30.3416
16.0359 | -35.2407 | -30.3503
16.042 | -35.2414 | -30.3588
16.0481 | -35.2421 | -30.3676
16.0542 | -35.2429 | -30.3764
16.0603 | -35.2436 | -30.3848
16.0664 | -35.2443 | -30.3934
16.0725 | -35.2451 | -30.402
16.0786 | -35.2459 | -30.4106
16.0847 | -35.2466 | -30.4193
16.0908 | -35.2473 | -30.4279
16.0969 | -35.248 | -30.4367
16.103 | -35.2488 | -30.4454
16.1091 | -35.2495 | -30.4541
16.1152 | -35.2503 | -30.4628
16.1214 | -35.251 | -30.4715
16.1275 | -35.2518 | -30.4802
16.1336 | -35.2525 | -30.4887
16.1397 | -35.2533 | -30.4974
16.1458 | -35.254 | -30.5059
16.1519 | -35.2548 | -30.5147
16.158 | -35.2555 | -30.5234
16.1641 | -35.2562 | -30.5323
16.1702 | -35.257 | -30.541
16.1763 | -35.2577 | -30.5497
16.1824 | -35.2585 | -30.5583
16.1885 | -35.2592 | -30.5667
16.1946 | -35.26 | -30.5754
16.2007 | -35.2607 | -30.5839
16.2068 | -35.2615 | -30.5926
16.2129 | -35.2622 | -30.6014
16.219 | -35.263 | -30.6096
16.2251 | -35.2637 | -30.6181
16.2312 | -35.2645 | -30.6267
16.2373 | -35.2652 | -30.6352
16.2434 | -35.2659 | -30.6438
16.2495 | -35.2667 | -30.6522
16.2556 | -35.2674 | -30.661
16.2618 | -35.2682 | -30.6695
16.2679 | -35.2689 | -30.6783
16.274 | -35.2696 | -30.6869
16.2801 | -35.2703 | -30.6954
16.2862 | -35.2711 | -30.7038
16.2923 | -35.2719 | -30.7125
16.2984 | -35.2726 | -30.7211
16.3045 | -35.2733 | -30.7297
16.3106 | -35.274 | -30.7384
16.3167 | -35.2748 | -30.7468
16.3228 | -35.2755 | -30.7553
16.3289 | -35.2763 | -30.7641
16.335 | -35.277 | -30.7725
16.3411 | -35.2777 | -30.7811
16.3472 | -35.2785 | -30.7899
16.3533 | -35.2793 | -30.7984
16.3594 | -35.28 | -30.8068
16.3655 | -35.2807 | -30.8154
16.3716 | -35.2814 | -30.8239
16.3777 | -35.2822 | -30.8325
16.3838 | -35.2829 | -30.841
16.3899 | -35.2836 | -30.8493
16.396 | -35.2844 | -30.8579
16.4021 | -35.2851 | -30.8663
16.4083 | -35.2859 | -30.8747
16.4144 | -35.2866 | -30.8834
16.4205 | -35.2874 | -30.892
16.4266 | -35.2881 | -30.9001
16.4327 | -35.2889 | -30.9089
16.4388 | -35.2896 | -30.9175
16.4449 | -35.2904 | -30.9261
16.451 | -35.2911 | -30.9346
16.4571 | -35.2919 | -30.9431
16.4632 | -35.2926 | -30.9515
16.4693 | -35.2933 | -30.96
16.4754 | -35.294 | -30.9687
16.4815 | -35.2948 | -30.9773
16.4876 | -35.2955 | -30.9855
16.4937 | -35.2963 | -30.9941
16.4998 | -35.297 | -31.0022
16.5059 | -35.2978 | -31.0106
16.512 | -35.2986 | -31.0194
16.5181 | -35.2993 | -31.0281
16.5242 | -35.3 | -31.0369
16.5303 | -35.3008 | -31.0452
16.5364 | -35.3015 | -31.0538
16.5425 | -35.3023 | -31.0621
16.5487 | -35.3029 | -31.0702
16.5548 | -35.3036 | -31.0781
16.5609 | -35.3043 | -31.0875
16.567 | -35.3051 | -31.0958
16.5731 | -35.3059 | -31.1045
16.5792 | -35.3066 | -31.1131
16.5853 | -35.3074 | -31.1218
16.5914 | -35.3081 | -31.1305
16.5975 | -35.3088 | -31.1392
16.6036 | -35.3095 | -31.1475
16.6097 | -35.3102 | -31.1562
16.6158 | -35.311 | -31.1648
16.6219 | -35.3117 | -31.1733
16.628 | -35.3124 | -31.1821
16.6341 | -35.3132 | -31.1905
16.6402 | -35.3139 | -31.1988
16.6463 | -35.3146 | -31.2074
16.6524 | -35.3153 | -31.2159
16.6585 | -35.316 | -31.2245
16.6646 | -35.3168 | -31.233
16.6707 | -35.3175 | -31.2411
16.6768 | -35.3182 | -31.2499
16.6829 | -35.319 | -31.2581
16.689 | -35.3197 | -31.2668
16.6952 | -35.3205 | -31.2751
16.7013 | -35.3212 | -31.2836
16.7074 | -35.322 | -31.2917
16.7135 | -35.3227 | -31.3
16.7196 | -35.3234 | -31.3085
16.7257 | -35.3242 | -31.3169
16.7318 | -35.3249 | -31.3254
16.7379 | -35.3256 | -31.3339
16.744 | -35.3264 | -31.3422
16.7501 | -35.3271 | -31.3507
16.7562 | -35.3278 | -31.3594
16.7623 | -35.3285 | -31.3679
16.7684 | -35.3293 | -31.3764
16.7745 | -35.33 | -31.3851
16.7806 | -35.3307 | -31.3933
16.7867 | -35.3315 | -31.4018
16.7928 | -35.3322 | -31.4102
16.7989 | -35.333 | -31.4185
16.805 | -35.3337 | -31.4268
16.8111 | -35.3344 | -31.4352
16.8172 | -35.3351 | -31.4436
16.8233 | -35.3359 | -31.4522
16.8294 | -35.3367 | -31.4607
16.8356 | -35.3374 | -31.4693
16.8417 | -35.3381 | -31.4777
16.8478 | -35.3389 | -31.486
16.8539 | -35.3396 | -31.4944
16.86 | -35.3403 | -31.503
16.8661 | -35.3411 | -31.5114
16.8722 | -35.3418 | -31.5199
16.8783 | -35.3425 | -31.5283
16.8844 | -35.3432 | -31.5365
16.8905 | -35.344 | -31.5449
16.8966 | -35.3447 | -31.5533
16.9027 | -35.3454 | -31.5618
16.9088 | -35.3461 | -31.5704
16.9149 | -35.3469 | -31.5788
16.921 | -35.3476 | -31.5868
16.9271 | -35.3483 | -31.5953
16.9332 | -35.349 | -31.6038
16.9393 | -35.3498 | -31.6123
16.9454 | -35.3505 | -31.6207
16.9515 | -35.3512 | -31.6291
16.9576 | -35.3519 | -31.6374
16.9637 | -35.3527 | -31.6456
16.9698 | -35.3534 | -31.654
16.9759 | -35.3541 | -31.6625
16.9821 | -35.3549 | -31.671
16.9882 | -35.3556 | -31.6791
16.9943 | -35.3563 | -31.6875
17.0004 | -35.3571 | -31.6961
17.0065 | -35.3578 | -31.7046
17.0126 | -35.3585 | -31.713
17.0187 | -35.3592 | -31.7214
17.0248 | -35.3599 | -31.7295
17.0309 | -35.3607 | -31.7379
17.037 | -35.3614 | -31.7463
17.0431 | -35.3621 | -31.7547
17.0492 | -35.3628 | -31.7631
17.0553 | -35.3636 | -31.7716
17.0614 | -35.3643 | -31.78
17.0675 | -35.365 | -31.7884
17.0736 | -35.3657 | -31.7969
17.0797 | -35.3665 | -31.8051
17.0858 | -35.3672 | -31.8135
17.0919 | -35.3679 | -31.8218
17.098 | -35.3686 | -31.8299
17.1041 | -35.3694 | -31.8383
17.1102 | -35.3701 | -31.8466
17.1163 | -35.3708 | -31.8549
17.1225 | -35.3715 | -31.8632
17.1286 | -35.3723 | -31.8712
17.1347 | -35.373 | -31.8795
17.1408 | -35.3737 | -31.888
17.1469 | -35.3744 | -31.8964
17.153 | -35.3752 | -31.905
17.1591 | -35.3759 | -31.9135
17.1652 | -35.3766 | -31.9216
17.1713 | -35.3774 | -31.9298
17.1774 | -35.3781 | -31.9378
17.1835 | -35.3789 | -31.9464
17.1896 | -35.3797 | -31.9549
17.1957 | -35.3804 | -31.9634
17.2018 | -35.3811 | -31.9712
17.2079 | -35.3819 | -31.9797
17.214 | -35.3826 | -31.9882
17.2201 | -35.3833 | -31.9965
17.2262 | -35.384 | -32.0051
17.2323 | -35.3847 | -32.0137
17.2384 | -35.3855 | -32.0218
17.2445 | -35.3862 | -32.03
17.2506 | -35.3869 | -32.0385
17.2567 | -35.3877 | -32.0468
17.2628 | -35.3884 | -32.055
17.269 | -35.3892 | -32.0631
17.2751 | -35.3899 | -32.0714
17.2812 | -35.3906 | -32.0799
17.2873 | -35.3913 | -32.0883
17.2934 | -35.3921 | -32.0968
17.2995 | -35.3928 | -32.1051
17.3056 | -35.3936 | -32.1132
17.3117 | -35.3943 | -32.1216
17.3178 | -35.395 | -32.1298
17.3239 | -35.3958 | -32.1383
17.33 | -35.3965 | -32.1467
17.3361 | -35.3972 | -32.1551
17.3422 | -35.398 | -32.1636
17.3483 | -35.3988 | -32.1717
17.3544 | -35.3995 | -32.1805
17.3605 | -35.4002 | -32.1891
17.3666 | -35.401 | -32.1974
17.3727 | -35.4017 | -32.2057
17.3788 | -35.4024 | -32.2137
17.3849 | -35.4031 | -32.2221
17.391 | -35.4038 | -32.2303
17.3971 | -35.4046 | -32.2385
17.4032 | -35.4053 | -32.2467
17.4094 | -35.406 | -32.2548
17.4155 | -35.4068 | -32.2631
17.4216 | -35.4075 | -32.2712
17.4277 | -35.4082 | -32.2795
17.4338 | -35.409 | -32.2878
17.4399 | -35.4097 | -32.2964
17.446 | -35.4104 | -32.3045
17.4521 | -35.4111 | -32.3127
17.4582 | -35.4118 | -32.3209
17.4643 | -35.4126 | -32.3291
17.4704 | -35.4133 | -32.3373
17.4765 | -35.414 | -32.3456
17.4826 | -35.4148 | -32.3537
17.4887 | -35.4155 | -32.3622
17.4948 | -35.4162 | -32.3707
17.5009 | -35.4169 | -32.3791
17.507 | -35.4176 | -32.3874
17.5131 | -35.4184 | -32.3957
17.5192 | -35.419 | -32.4035
17.5253 | -35.4198 | -32.4119
17.5314 | -35.4205 | -32.4203
17.5375 | -35.4212 | -32.4288
17.5436 | -35.4219 | -32.4372
17.5497 | -35.4227 | -32.4455
17.5559 | -35.4234 | -32.4536
17.562 | -35.4241 | -32.4619
17.5681 | -35.4248 | -32.4699
17.5742 | -35.4256 | -32.4782
17.5803 | -35.4263 | -32.4864
17.5864 | -35.4271 | -32.4945
17.5925 | -35.4278 | -32.5028
17.5986 | -35.4285 | -32.5111
17.6047 | -35.4293 | -32.5194
17.6108 | -35.43 | -32.5276
17.6169 | -35.4307 | -32.5358
17.623 | -35.4314 | -32.5436
17.6291 | -35.4322 | -32.5518
17.6352 | -35.4329 | -32.56
17.6413 | -35.4336 | -32.5683
17.6474 | -35.4343 | -32.5767
17.6535 | -35.4351 | -32.5851
17.6596 | -35.4358 | -32.593
17.6657 | -35.4366 | -32.6014
17.6718 | -35.4373 | -32.6095
17.6779 | -35.438 | -32.6178
17.684 | -35.4387 | -32.6261
17.6901 | -35.4394 | -32.6345
17.6963 | -35.4402 | -32.6425
17.7024 | -35.4409 | -32.6508
17.7085 | -35.4416 | -32.6589
17.7146 | -35.4423 | -32.6672
17.7207 | -35.4431 | -32.6753
17.7268 | -35.4438 | -32.6833
17.7329 | -35.4445 | -32.6914
17.739 | -35.4453 | -32.6996
17.7451 | -35.446 | -32.7077
17.7512 | -35.4468 | -32.7159
17.7573 | -35.4475 | -32.7244
17.7634 | -35.4482 | -32.7322
17.7695 | -35.449 | -32.7404
17.7756 | -35.4497 | -32.7489
17.7817 | -35.4504 | -32.7568
17.7878 | -35.4511 | -32.7651
17.7939 | -35.4518 | -32.7733
17.8 | -35.4525 | -32.7812
17.8061 | -35.4533 | -32.7894
17.8122 | -35.454 | -32.7977
17.8183 | -35.4548 | -32.8058
17.8244 | -35.4555 | -32.8138
17.8305 | -35.4562 | -32.8221
17.8366 | -35.4569 | -32.8299
17.8428 | -35.4577 | -32.8383
17.8489 | -35.4584 | -32.8464
17.855 | -35.4591 | -32.8546
17.8611 | -35.4598 | -32.8627
17.8672 | -35.4606 | -32.8707
17.8733 | -35.4613 | -32.8789
17.8794 | -35.4621 | -32.8872
17.8855 | -35.4628 | -32.8954
17.8916 | -35.4635 | -32.9033
17.8977 | -35.4642 | -32.9116
17.9038 | -35.4649 | -32.9196
17.9099 | -35.4656 | -32.9278
17.916 | -35.4664 | -32.9362
17.9221 | -35.4672 | -32.944
17.9282 | -35.468 | -32.9515
17.9343 | -35.4686 | -32.959
17.9404 | -35.4695 | -32.9669
17.9465 | -35.4703 | -32.9751
17.9526 | -35.4709 | -32.9837
17.9587 | -35.4717 | -32.9919
17.9648 | -35.4724 | -32.9999
17.9709 | -35.4732 | -33.008
17.977 | -35.4739 | -33.0156
17.9832 | -35.4747 | -33.0238
17.9893 | -35.4754 | -33.0322
17.9954 | -35.4761 | -33.0405
18.0015 | -35.4768 | -33.0486
18.0076 | -35.4776 | -33.0564
18.0137 | -35.4783 | -33.0645
18.0198 | -35.479 | -33.0725
18.0259 | -35.4798 | -33.0805
18.032 | -35.4805 | -33.0885
18.0381 | -35.4812 | -33.097
18.0442 | -35.4819 | -33.1046
18.0503 | -35.4827 | -33.1128
18.0564 | -35.4834 | -33.121
18.0625 | -35.4841 | -33.1293
18.0686 | -35.4849 | -33.1377
18.0747 | -35.4856 | -33.1457
18.0808 | -35.4863 | -33.1531
18.0869 | -35.4871 | -33.1611
18.093 | -35.4878 | -33.169
18.0991 | -35.4885 | -33.1772
18.1052 | -35.4892 | -33.1851
18.1113 | -35.49 | -33.1934
18.1174 | -35.4907 | -33.2006
18.1236 | -35.4914 | -33.2089
18.1297 | -35.4921 | -33.2173
18.1358 | -35.4928 | -33.2249
18.1419 | -35.4935 | -33.2331
18.148 | -35.4942 | -33.2409
18.1541 | -35.4949 | -33.2486
18.1602 | -35.4956 | -33.2566
18.1663 | -35.4964 | -33.2644
18.1724 | -35.4971 | -33.2725
18.1785 | -35.4978 | -33.2803
18.1846 | -35.4985 | -33.2881
18.1907 | -35.4992 | -33.296
18.1968 | -35.4999 | -33.3042
18.2029 | -35.5007 | -33.3122
18.209 | -35.5015 | -33.3204
18.2151 | -35.5022 | -33.3285
18.2212 | -35.5029 | -33.3364
18.2273 | -35.5037 | -33.3443
18.2334 | -35.5044 | -33.3524
18.2395 | -35.5051 | -33.3605
18.2456 | -35.5059 | -33.3686
18.2517 | -35.5066 | -33.3764
18.2578 | -35.5073 | -33.3841
18.2639 | -35.5081 | -33.3921
18.2701 | -35.5088 | -33.4003
18.2762 | -35.5096 | -33.408
18.2823 | -35.5103 | -33.4161
18.2884 | -35.511 | -33.4239
18.2945 | -35.5118 | -33.4316
18.3006 | -35.5125 | -33.4398
18.3067 | -35.5133 | -33.4477
18.3128 | -35.514 | -33.456
18.3189 | -35.5147 | -33.4638
18.325 | -35.5154 | -33.4715
18.3311 | -35.5161 | -33.4791
18.3372 | -35.5168 | -33.4871
18.3433 | -35.5175 | -33.4949
18.3494 | -35.5182 | -33.5031
18.3555 | -35.519 | -33.5109
18.3616 | -35.5197 | -33.5185
18.3677 | -35.5204 | -33.5266
18.3738 | -35.5212 | -33.5348
18.3799 | -35.5219 | -33.5429
18.386 | -35.5226 | -33.5509
18.3921 | -35.5233 | -33.5589
18.3982 | -35.524 | -33.5662
18.4043 | -35.5247 | -33.5744
18.4105 | -35.5254 | -33.5821
18.4166 | -35.5261 | -33.59
18.4227 | -35.5268 | -33.598
18.4288 | -35.5276 | -33.6061
18.4349 | -35.5283 | -33.6138
18.441 | -35.529 | -33.6215
18.4471 | -35.5297 | -33.6294
18.4532 | -35.5304 | -33.6373
18.4593 | -35.5311 | -33.6454
18.4654 | -35.5318 | -33.6528
18.4715 | -35.5326 | -33.6607
18.4776 | -35.5333 | -33.6689
18.4837 | -35.534 | -33.677
18.4898 | -35.5347 | -33.6851
18.4959 | -35.5354 | -33.6931
18.502 | -35.5362 | -33.7011
18.5081 | -35.5369 | -33.7092
18.5142 | -35.5376 | -33.717
18.5203 | -35.5383 | -33.725
18.5264 | -35.5389 | -33.7328
18.5325 | -35.5397 | -33.7406
18.5386 | -35.5404 | -33.7486
18.5447 | -35.5411 | -33.7566
18.5508 | -35.5419 | -33.7645
18.557 | -35.5426 | -33.7724
18.5631 | -35.5432 | -33.7804
18.5692 | -35.544 | -33.7883
18.5753 | -35.5446 | -33.7958
18.5814 | -35.5453 | -33.8035
18.5875 | -35.5461 | -33.8115
18.5936 | -35.5468 | -33.8193
18.5997 | -35.5475 | -33.8273
18.6058 | -35.5482 | -33.835
18.6119 | -35.5489 | -33.8428
18.618 | -35.5496 | -33.8508
18.6241 | -35.5503 | -33.8586
18.6302 | -35.5511 | -33.8666
18.6363 | -35.5518 | -33.8744
18.6424 | -35.5525 | -33.882
18.6485 | -35.5532 | -33.8898
18.6546 | -35.5539 | -33.8978
18.6607 | -35.5547 | -33.9056
18.6668 | -35.5554 | -33.9136
18.6729 | -35.5561 | -33.9219
18.679 | -35.5569 | -33.9296
18.6851 | -35.5576 | -33.9372
18.6912 | -35.5583 | -33.9451
18.6974 | -35.559 | -33.9532
18.7035 | -35.5597 | -33.961
18.7096 | -35.5604 | -33.9691
18.7157 | -35.5612 | -33.9769
18.7218 | -35.5619 | -33.9847
18.7279 | -35.5626 | -33.9926
18.734 | -35.5634 | -34.0008
18.7401 | -35.5641 | -34.0088
18.7462 | -35.5648 | -34.0163
18.7523 | -35.5655 | -34.0242
18.7584 | -35.5663 | -34.0319
18.7645 | -35.567 | -34.04
18.7706 | -35.5677 | -34.0477
18.7767 | -35.5684 | -34.0557
18.7828 | -35.5691 | -34.0631
18.7889 | -35.5698 | -34.071
18.795 | -35.5706 | -34.0789
18.8011 | -35.5712 | -34.0868
18.8072 | -35.572 | -34.0945
18.8133 | -35.5727 | -34.1023
18.8194 | -35.5734 | -34.1097
18.8255 | -35.5741 | -34.1176
18.8316 | -35.5748 | -34.1254
18.8377 | -35.5755 | -34.1331
18.8439 | -35.5763 | -34.1409
18.85 | -35.577 | -34.1489
18.8561 | -35.5777 | -34.1563
18.8622 | -35.5784 | -34.1641
18.8683 | -35.5791 | -34.1718
18.8744 | -35.5798 | -34.1796
18.8805 | -35.5805 | -34.1878
18.8866 | -35.5812 | -34.1956
18.8927 | -35.5819 | -34.2033
18.8988 | -35.5827 | -34.2111
18.9049 | -35.5834 | -34.2192
18.911 | -35.5841 | -34.2269
18.9171 | -35.5848 | -34.2348
18.9232 | -35.5857 | -34.242
18.9293 | -35.5866 | -34.2499
18.9354 | -35.5872 | -34.2571
18.9415 | -35.5879 | -34.2649
18.9476 | -35.5885 | -34.2723
18.9537 | -35.5893 | -34.28
18.9598 | -35.5899 | -34.2877
18.9659 | -35.5906 | -34.2956
18.972 | -35.5914 | -34.3033
18.9781 | -35.5921 | -34.311
18.9843 | -35.5928 | -34.3187
18.9904 | -35.5935 | -34.3263
18.9965 | -35.5942 | -34.3337
19.0026 | -35.5949 | -34.3416
19.0087 | -35.5957 | -34.3494
19.0148 | -35.5964 | -34.3572
19.0209 | -35.5971 | -34.3649
19.027 | -35.5978 | -34.3727
19.0331 | -35.5985 | -34.3799
19.0392 | -35.5992 | -34.3879
19.0453 | -35.6 | -34.3958
19.0514 | -35.6007 | -34.4035
19.0575 | -35.6014 | -34.4115
19.0636 | -35.6022 | -34.4186
19.0697 | -35.6029 | -34.4263
19.0758 | -35.6036 | -34.434
19.0819 | -35.6043 | -34.4418
19.088 | -35.6051 | -34.4496
19.0941 | -35.6058 | -34.4574
19.1002 | -35.6064 | -34.465
19.1063 | -35.6071 | -34.4728
19.1124 | -35.6078 | -34.4802
19.1185 | -35.6085 | -34.4881
19.1246 | -35.6092 | -34.4956
19.1308 | -35.61 | -34.5038
19.1369 | -35.6107 | -34.5111
19.143 | -35.6114 | -34.5187
19.1491 | -35.6121 | -34.5263
19.1552 | -35.6128 | -34.5342
19.1613 | -35.6135 | -34.5419
19.1674 | -35.6142 | -34.5498
19.1735 | -35.6149 | -34.5566
19.1796 | -35.6156 | -34.5638
19.1857 | -35.6163 | -34.5717
19.1918 | -35.617 | -34.5789
19.1979 | -35.6177 | -34.5862
19.204 | -35.6184 | -34.594
19.2101 | -35.6192 | -34.6013
19.2162 | -35.6199 | -34.6092
19.2223 | -35.6206 | -34.6171
19.2284 | -35.6212 | -34.6243
19.2345 | -35.6219 | -34.6317
19.2406 | -35.6226 | -34.6392
19.2467 | -35.6233 | -34.6466
19.2528 | -35.624 | -34.6541
19.2589 | -35.6247 | -34.661
19.265 | -35.6254 | -34.6687
19.2712 | -35.6261 | -34.6763
19.2773 | -35.6268 | -34.6839
19.2834 | -35.6275 | -34.6917
19.2895 | -35.6282 | -34.6992
19.2956 | -35.6289 | -34.7066
19.3017 | -35.6296 | -34.7138
19.3078 | -35.6303 | -34.7216
19.3139 | -35.631 | -34.7283
19.32 | -35.6317 | -34.7357
19.3261 | -35.6324 | -34.7432
19.3322 | -35.6331 | -34.7508
19.3383 | -35.6337 | -34.7581
19.3444 | -35.6344 | -34.7655
19.3505 | -35.6351 | -34.7731
19.3566 | -35.6358 | -34.7808
19.3627 | -35.6365 | -34.7881
19.3688 | -35.6372 | -34.7957
19.3749 | -35.638 | -34.8036
19.381 | -35.6387 | -34.8111
19.3871 | -35.6394 | -34.8184
19.3932 | -35.6401 | -34.8256
19.3993 | -35.6409 | -34.8331
19.4054 | -35.6416 | -34.8404
19.4115 | -35.6422 | -34.8484
19.4177 | -35.6428 | -34.8554
19.4238 | -35.6435 | -34.8644
19.4299 | -35.644 | -34.8713
19.436 | -35.6445 | -34.8803
19.4421 | -35.6451 | -34.8888
19.4482 | -35.6458 | -34.8955
19.4543 | -35.6465 | -34.9029
19.4604 | -35.6472 | -34.9106
19.4665 | -35.6479 | -34.918
19.4726 | -35.6486 | -34.9257
19.4787 | -35.6492 | -34.9328
19.4848 | -35.6499 | -34.9403
19.4909 | -35.6506 | -34.9474
19.497 | -35.6513 | -34.9551
19.5031 | -35.6519 | -34.9624
19.5092 | -35.6527 | -34.9699
19.5153 | -35.6533 | -34.9774
19.5214 | -35.654 | -34.9845
19.5275 | -35.6547 | -34.9919
19.5336 | -35.6553 | -34.9999
19.5397 | -35.656 | -35.0071
19.5458 | -35.6567 | -35.0146
19.5519 | -35.6574 | -35.0219
19.5581 | -35.6581 | -35.0289
19.5642 | -35.6587 | -35.0364
19.5703 | -35.6593 | -35.0441
19.5764 | -35.66 | -35.0516
19.5825 | -35.6607 | -35.0593
19.5886 | -35.6614 | -35.0668
19.5947 | -35.6621 | -35.0742
19.6008 | -35.6627 | -35.0818
19.6069 | -35.6635 | -35.0893
19.613 | -35.6642 | -35.0968
19.6191 | -35.6648 | -35.1041
19.6252 | -35.6655 | -35.1115
19.6313 | -35.6661 | -35.1189
19.6374 | -35.6668 | -35.1261
19.6435 | -35.6675 | -35.1334
19.6496 | -35.6681 | -35.141
19.6557 | -35.6688 | -35.1487
19.6618 | -35.6695 | -35.1553
19.6679 | -35.6701 | -35.1624
19.674 | -35.6708 | -35.1697
19.6801 | -35.6715 | -35.1774
19.6862 | -35.6722 | -35.1848
19.6923 | -35.673 | -35.1918
19.6984 | -35.6735 | -35.1989
19.7046 | -35.6743 | -35.2064
19.7107 | -35.6749 | -35.2136
19.7168 | -35.6756 | -35.2212
19.7229 | -35.6762 | -35.2287
19.729 | -35.6769 | -35.2362
19.7351 | -35.6775 | -35.2426
19.7412 | -35.6782 | -35.2507
19.7473 | -35.679 | -35.2577
19.7534 | -35.6796 | -35.2649
19.7595 | -35.6802 | -35.2721
19.7656 | -35.6809 | -35.2796
19.7717 | -35.6815 | -35.2865
19.7778 | -35.6822 | -35.2942
19.7839 | -35.6829 | -35.3011
19.79 | -35.6835 | -35.3089
19.7961 | -35.6842 | -35.3164
19.8022 | -35.6848 | -35.3233
19.8083 | -35.6854 | -35.3304
19.8144 | -35.6861 | -35.3375
19.8205 | -35.6867 | -35.3451
19.8266 | -35.6874 | -35.3525
19.8327 | -35.6881 | -35.3601
19.8388 | -35.6887 | -35.3667
19.845 | -35.6893 | -35.374
19.8511 | -35.69 | -35.3809
19.8572 | -35.6907 | -35.3882
19.8633 | -35.6914 | -35.3958
19.8694 | -35.6921 | -35.4032
19.8755 | -35.6928 | -35.4098
19.8816 | -35.6935 | -35.417
19.8877 | -35.6942 | -35.424
19.8938 | -35.6949 | -35.4312
19.8999 | -35.6955 | -35.4391
19.906 | -35.6962 | -35.4465
19.9121 | -35.6968 | -35.4538
19.9182 | -35.6975 | -35.4614
19.9243 | -35.6982 | -35.4688
19.9304 | -35.6988 | -35.4757
19.9365 | -35.6995 | -35.4829
19.9426 | -35.7003 | -35.49
19.9487 | -35.701 | -35.4969
19.9548 | -35.7016 | -35.5043
19.9609 | -35.7022 | -35.5119
19.967 | -35.7029 | -35.5189
19.9731 | -35.7036 | -35.5258
19.9792 | -35.7044 | -35.5326
19.9853 | -35.705 | -35.5397
19.9915 | -35.7057 | -35.547
19.9976 | -35.7063 | -35.5541
20.0037 | -35.7069 | -35.5609
20.0098 | -35.7076 | -35.5683
20.0159 | -35.7082 | -35.5754
20.022 | -35.7088 | -35.5829
20.0281 | -35.7095 | -35.5902
20.0342 | -35.7101 | -35.5977
20.0403 | -35.7108 | -35.6044
20.0464 | -35.7115 | -35.612
20.0525 | -35.7121 | -35.6187
20.0586 | -35.7127 | -35.6259
20.0647 | -35.7134 | -35.6333
20.0708 | -35.7141 | -35.6405
20.0769 | -35.7148 | -35.6478
20.083 | -35.7155 | -35.6548
20.0891 | -35.7161 | -35.6621
20.0952 | -35.7167 | -35.6691
20.1013 | -35.7174 | -35.6761
20.1074 | -35.7181 | -35.6831
20.1135 | -35.7187 | -35.6902
20.1196 | -35.7194 | -35.6975
20.1257 | -35.72 | -35.7044
20.1319 | -35.7206 | -35.7115
20.138 | -35.7213 | -35.7187
20.1441 | -35.7219 | -35.7259
20.1502 | -35.7226 | -35.7329
20.1563 | -35.7232 | -35.7392
20.1624 | -35.724 | -35.7475
20.1685 | -35.7248 | -35.7559
20.1746 | -35.7254 | -35.765
20.1807 | -35.7263 | -35.7712
20.1868 | -35.7269 | -35.7798
20.1929 | -35.7274 | -35.7863
20.199 | -35.728 | -35.7937
20.2051 | -35.7288 | -35.8008
20.2112 | -35.7293 | -35.8078
20.2173 | -35.73 | -35.8145
20.2234 | -35.7307 | -35.8218
20.2295 | -35.7313 | -35.8284
20.2356 | -35.732 | -35.8354
20.2417 | -35.7326 | -35.8431
20.2478 | -35.7333 | -35.8505
20.2539 | -35.7339 | -35.8577
20.26 | -35.7346 | -35.8646
20.2661 | -35.7352 | -35.8717
20.2723 | -35.7359 | -35.8787
20.2784 | -35.7365 | -35.8859
20.2845 | -35.7372 | -35.8933
20.2906 | -35.7378 | -35.9005
20.2967 | -35.7384 | -35.9078
20.3028 | -35.7391 | -35.9146
20.3089 | -35.7398 | -35.9216
20.315 | -35.7404 | -35.9289
20.3211 | -35.7411 | -35.9361
20.3272 | -35.7417 | -35.9436
20.3333 | -35.7424 | -35.9505
20.3394 | -35.743 | -35.9576
20.3455 | -35.7437 | -35.9648
20.3516 | -35.7443 | -35.9719
20.3577 | -35.7449 | -35.979
20.3638 | -35.7456 | -35.9864
20.3699 | -35.7462 | -35.993
20.376 | -35.7467 | -36
20.3821 | -35.7474 | -36.0073
20.3882 | -35.7479 | -36.0143
20.3943 | -35.7486 | -36.0215
20.4004 | -35.7493 | -36.0281
20.4065 | -35.7499 | -36.0353
20.4126 | -35.7506 | -36.0424
20.4188 | -35.7512 | -36.0495
20.4249 | -35.7519 | -36.0562
20.431 | -35.7525 | -36.0633
20.4371 | -35.7532 | -36.0697
20.4432 | -35.7538 | -36.0768
20.4493 | -35.7544 | -36.0837
20.4554 | -35.7551 | -36.0905
20.4615 | -35.7558 | -36.0973
20.4676 | -35.7564 | -36.104
20.4737 | -35.7571 | -36.1107
20.4798 | -35.7577 | -36.1181
20.4859 | -35.7583 | -36.1257
20.492 | -35.7588 | -36.1335
20.4981 | -35.7597 | -36.1409
20.5042 | -35.76 | -36.1485
20.5103 | -35.7604 | -36.1578
20.5164 | -35.7609 | -36.1665
20.5225 | -35.7615 | -36.1725
20.5286 | -35.7621 | -36.1796
20.5347 | -35.7628 | -36.1865
20.5408 | -35.7635 | -36.1934
20.5469 | -35.7641 | -36.2003
20.553 | -35.7648 | -36.2077
20.5592 | -35.7655 | -36.2142
20.5653 | -35.7661 | -36.2212
20.5714 | -35.7667 | -36.2283
20.5775 | -35.7673 | -36.2354
20.5836 | -35.7679 | -36.2428
20.5897 | -35.7685 | -36.2499
20.5958 | -35.7691 | -36.2575
20.6019 | -35.7697 | -36.2646
20.608 | -35.7703 | -36.2715
20.6141 | -35.7709 | -36.2786
20.6202 | -35.7716 | -36.2853
20.6263 | -35.7723 | -36.292
20.6324 | -35.7729 | -36.2994
20.6385 | -35.7735 | -36.3063
20.6446 | -35.7741 | -36.3135
20.6507 | -35.7747 | -36.3207
20.6568 | -35.7754 | -36.3278
20.6629 | -35.776 | -36.3353
20.669 | -35.7766 | -36.3422
20.6751 | -35.7773 | -36.3494
20.6812 | -35.7779 | -36.3556
20.6873 | -35.7785 | -36.3629
20.6934 | -35.7791 | -36.3699
20.6995 | -35.7798 | -36.3769
20.7057 | -35.7804 | -36.3841
20.7118 | -35.7811 | -36.3913
20.7179 | -35.7817 | -36.3977
20.724 | -35.7824 | -36.4047
20.7301 | -35.783 | -36.4117
20.7362 | -35.7837 | -36.4187
20.7423 | -35.7843 | -36.4256
20.7484 | -35.785 | -36.4328
20.7545 | -35.7857 | -36.4396
20.7606 | -35.7863 | -36.4468
20.7667 | -35.787 | -36.4542
20.7728 | -35.7876 | -36.4614
20.7789 | -35.7882 | -36.4688
20.785 | -35.7889 | -36.4762
20.7911 | -35.7896 | -36.4829
20.7972 | -35.7902 | -36.4897
20.8033 | -35.7908 | -36.497
20.8094 | -35.7915 | -36.5041
20.8155 | -35.7921 | -36.511
20.8216 | -35.7928 | -36.5178
20.8277 | -35.7934 | -36.5242
20.8338 | -35.794 | -36.5318
20.8399 | -35.7945 | -36.5396
20.8461 | -35.7953 | -36.5475
20.8522 | -35.796 | -36.555
20.8583 | -35.7966 | -36.5603
20.8644 | -35.7972 | -36.5679
20.8705 | -35.7979 | -36.575
20.8766 | -35.7983 | -36.5821
20.8827 | -35.7989 | -36.5883
20.8888 | -35.7996 | -36.5957
20.8949 | -35.8004 | -36.6025
20.901 | -35.8012 | -36.6093
20.9071 | -35.8018 | -36.6149
20.9132 | -35.8023 | -36.6226
20.9193 | -35.803 | -36.6311
20.9254 | -35.8036 | -36.6396
20.9315 | -35.8044 | -36.6465
20.9376 | -35.805 | -36.6539
20.9437 | -35.8057 | -36.6609
20.9498 | -35.8063 | -36.6675
20.9559 | -35.8068 | -36.6746
20.962 | -35.8075 | -36.6811
20.9681 | -35.8081 | -36.6887
20.9742 | -35.8087 | -36.6958
20.9803 | -35.8093 | -36.7027
20.9864 | -35.8099 | -36.7098
20.9926 | -35.8106 | -36.7169
20.9987 | -35.8112 | -36.7232
21.0048 | -35.8119 | -36.73
21.0109 | -35.8125 | -36.7365
21.017 | -35.8131 | -36.7433
21.0231 | -35.8138 | -36.7505
21.0292 | -35.8144 | -36.7578
21.0353 | -35.8149 | -36.7646
21.0414 | -35.8154 | -36.7714
21.0475 | -35.8161 | -36.7786
21.0536 | -35.8167 | -36.7853
21.0597 | -35.8173 | -36.7921
21.0658 | -35.8179 | -36.7993
21.0719 | -35.8185 | -36.8058
21.078 | -35.8191 | -36.813
21.0841 | -35.8197 | -36.8201
21.0902 | -35.8203 | -36.8272
21.0963 | -35.8209 | -36.8348
21.1024 | -35.8215 | -36.8421
21.1085 | -35.8222 | -36.8489
21.1146 | -35.8228 | -36.8557
21.1207 | -35.8235 | -36.8623
21.1268 | -35.8241 | -36.8695
21.133 | -35.8247 | -36.8766
21.1391 | -35.8254 | -36.8833
21.1452 | -35.826 | -36.8904
21.1513 | -35.8266 | -36.897
21.1574 | -35.8273 | -36.9043
21.1635 | -35.828 | -36.9119
21.1696 | -35.8286 | -36.9187
21.1757 | -35.8292 | -36.9258
21.1818 | -35.8299 | -36.9339
21.1879 | -35.8304 | -36.9415
21.194 | -35.831 | -36.9481
21.2001 | -35.8317 | -36.9558
21.2062 | -35.8323 | -36.9634
21.2123 | -35.833 | -36.9705
21.2184 | -35.8335 | -36.9773
21.2245 | -35.8341 | -36.9845
21.2306 | -35.8346 | -36.9917
21.2367 | -35.8352 | -36.9981
21.2428 | -35.8358 | -37.0049
21.2489 | -35.8365 | -37.0121
21.255 | -35.8371 | -37.0193
21.2611 | -35.8377 | -37.0264
21.2672 | -35.8384 | -37.0338
21.2733 | -35.8391 | -37.0387
21.2795 | -35.839 | -37.0416
21.2856 | -35.8392 | -37.0509
21.2917 | -35.8404 | -37.0588
21.2978 | -35.8412 | -37.0677
21.3039 | -35.8419 | -37.0753
21.31 | -35.8427 | -37.0814
21.3161 | -35.8435 | -37.0885
21.3222 | -35.8441 | -37.0954
21.3283 | -35.8447 | -37.1028
21.3344 | -35.8454 | -37.1098
21.3405 | -35.8461 | -37.1169
21.3466 | -35.8467 | -37.1244
21.3527 | -35.8473 | -37.1309
21.3588 | -35.8478 | -37.138
21.3649 | -35.8485 | -37.1453
21.371 | -35.8492 | -37.152
21.3771 | -35.8498 | -37.1592
21.3832 | -35.8504 | -37.1664
21.3893 | -35.8511 | -37.1728
21.3954 | -35.8518 | -37.1794
21.4015 | -35.8524 | -37.1861
21.4076 | -35.853 | -37.1938
21.4137 | -35.8537 | -37.2005
21.4199 | -35.8544 | -37.2076
21.426 | -35.8551 | -37.2148
21.4321 | -35.8558 | -37.2213
21.4382 | -35.8563 | -37.2285
21.4443 | -35.8569 | -37.2349
21.4504 | -35.8575 | -37.2419
21.4565 | -35.8581 | -37.2493
21.4626 | -35.8587 | -37.2558
21.4687 | -35.8593 | -37.2635
21.4748 | -35.86 | -37.271
21.4809 | -35.8606 | -37.2782
21.487 | -35.8612 | -37.2856
21.4931 | -35.8619 | -37.2927
21.4992 | -35.8625 | -37.3001
21.5053 | -35.8632 | -37.307
21.5114 | -35.8639 | -37.3144
21.5175 | -35.8644 | -37.3209
21.5236 | -35.8651 | -37.328
21.5297 | -35.8656 | -37.3353
21.5358 | -35.8662 | -37.342
21.5419 | -35.8667 | -37.3488
21.548 | -35.8673 | -37.3559
21.5541 | -35.8678 | -37.3636
21.5602 | -35.8684 | -37.3707
21.5664 | -35.8691 | -37.3777
21.5725 | -35.8697 | -37.3846
21.5786 | -35.8704 | -37.3912
21.5847 | -35.8709 | -37.3983
21.5908 | -35.8715 | -37.4053
21.5969 | -35.8722 | -37.4117
21.603 | -35.8727 | -37.419
21.6091 | -35.8733 | -37.4263
21.6152 | -35.874 | -37.4333
21.6213 | -35.8747 | -37.4403
21.6274 | -35.8753 | -37.4477
21.6335 | -35.8759 | -37.4534
21.6396 | -35.8765 | -37.4606
21.6457 | -35.8771 | -37.4685
21.6518 | -35.8779 | -37.4761
21.6579 | -35.8783 | -37.4833
21.664 | -35.8787 | -37.4896
21.6701 | -35.8792 | -37.4955
21.6762 | -35.8798 | -37.5029
21.6823 | -35.8807 | -37.5107
21.6884 | -35.8813 | -37.5168
21.6945 | -35.8821 | -37.5228
21.7006 | -35.8826 | -37.5308
21.7068 | -35.8832 | -37.537
21.7129 | -35.8839 | -37.5438
21.719 | -35.8844 | -37.5515
21.7251 | -35.8851 | -37.5576
21.7312 | -35.8856 | -37.5646
21.7373 | -35.8863 | -37.5703
21.7434 | -35.8869 | -37.5765
21.7495 | -35.8874 | -37.584
21.7556 | -35.888 | -37.5919
21.7617 | -35.8887 | -37.5995
21.7678 | -35.8893 | -37.606
21.7739 | -35.8901 | -37.6127
21.78 | -35.8907 | -37.6194
21.7861 | -35.8913 | -37.6267
21.7922 | -35.8919 | -37.6339
21.7983 | -35.8925 | -37.6403
21.8044 | -35.8932 | -37.6469
21.8105 | -35.8938 | -37.6537
21.8166 | -35.8944 | -37.6612
21.8227 | -35.895 | -37.6682
21.8288 | -35.8957 | -37.6751
21.8349 | -35.8963 | -37.6827
21.841 | -35.8969 | -37.6892
21.8471 | -35.8975 | -37.6957
21.8533 | -35.8981 | -37.7032
21.8594 | -35.8988 | -37.7099
21.8655 | -35.8995 | -37.7169
21.8716 | -35.9002 | -37.7239
21.8777 | -35.9008 | -37.7316
21.8838 | -35.9015 | -37.7382
21.8899 | -35.9021 | -37.7457
21.896 | -35.9027 | -37.7525
21.9021 | -35.9034 | -37.7593
21.9082 | -35.904 | -37.7662
21.9143 | -35.9046 | -37.7728
21.9204 | -35.9052 | -37.7796
21.9265 | -35.9058 | -37.7867
21.9326 | -35.9064 | -37.7933
21.9387 | -35.907 | -37.8003
21.9448 | -35.9076 | -37.807
21.9509 | -35.9082 | -37.8142
21.957 | -35.9088 | -37.8211
21.9631 | -35.9094 | -37.828
21.9692 | -35.9101 | -37.835
21.9753 | -35.9107 | -37.8422
21.9814 | -35.9113 | -37.8495
21.9875 | -35.912 | -37.8563
21.9937 | -35.9126 | -37.8635
21.9998 | -35.9132 | -37.8703
22.0059 | -35.9138 | -37.8772
22.012 | -35.9144 | -37.8838
22.0181 | -35.9151 | -37.89
22.0242 | -35.9157 | -37.8975
22.0303 | -35.9163 | -37.9048
22.0364 | -35.917 | -37.912
22.0425 | -35.9177 | -37.9186
22.0486 | -35.9183 | -37.9252
22.0547 | -35.9193 | -37.9308
22.0608 | -35.9198 | -37.9381
22.0669 | -35.9205 | -37.945
22.073 | -35.921 | -37.9511
22.0791 | -35.9213 | -37.958
22.0852 | -35.9218 | -37.9672
22.0913 | -35.9226 | -37.9748
22.0974 | -35.9232 | -37.9821
22.1035 | -35.9237 | -37.9891
22.1096 | -35.9244 | -37.9965
22.1157 | -35.9251 | -38.0044
22.1218 | -35.9257 | -38.0109
22.1279 | -35.9264 | -38.0178
22.134 | -35.927 | -38.0249
22.1402 | -35.9276 | -38.0322
22.1463 | -35.9282 | -38.0397
22.1524 | -35.9288 | -38.0465
22.1585 | -35.9295 | -38.0537
22.1646 | -35.9301 | -38.061
22.1707 | -35.9307 | -38.0682
22.1768 | -35.9313 | -38.0749
22.1829 | -35.9319 | -38.0818
22.189 | -35.9325 | -38.089
22.1951 | -35.9331 | -38.0962
22.2012 | -35.9338 | -38.103
22.2073 | -35.9345 | -38.1101
22.2134 | -35.9351 | -38.1172
22.2195 | -35.9358 | -38.1239
22.2256 | -35.9364 | -38.131
22.2317 | -35.9371 | -38.138
22.2378 | -35.9377 | -38.1455
22.2439 | -35.9384 | -38.1523
22.25 | -35.9391 | -38.1597
22.2561 | -35.9398 | -38.1668
22.2622 | -35.9404 | -38.1738
22.2683 | -35.941 | -38.181
22.2744 | -35.9417 | -38.1879
22.2806 | -35.9424 | -38.1948
22.2867 | -35.943 | -38.2022
22.2928 | -35.9437 | -38.2091
22.2989 | -35.9443 | -38.2166
22.305 | -35.945 | -38.2239
22.3111 | -35.9456 | -38.2307
22.3172 | -35.9463 | -38.2379
22.3233 | -35.947 | -38.2444
22.3294 | -35.9477 | -38.2514
22.3355 | -35.9482 | -38.258
22.3416 | -35.9488 | -38.2654
22.3477 | -35.9494 | -38.2725
22.3538 | -35.95 | -38.2795
22.3599 | -35.9506 | -38.2867
22.366 | -35.9512 | -38.2937
22.3721 | -35.9518 | -38.3007
22.3782 | -35.9524 | -38.3079
22.3843 | -35.953 | -38.3151
22.3904 | -35.9537 | -38.3218
22.3965 | -35.9542 | -38.3288
22.4026 | -35.9549 | -38.3349
22.4087 | -35.9557 | -38.3424
22.4148 | -35.9564 | -38.3483
22.4209 | -35.957 | -38.3542
22.4271 | -35.9575 | -38.3612
22.4332 | -35.9582 | -38.3679
22.4393 | -35.9588 | -38.3747
22.4454 | -35.9593 | -38.3817
22.4515 | -35.96 | -38.3882
22.4576 | -35.9606 | -38.3955
22.4637 | -35.9612 | -38.4028
22.4698 | -35.9619 | -38.4099
22.4759 | -35.9626 | -38.4174
22.482 | -35.9633 | -38.4249
22.4881 | -35.964 | -38.4321
22.4942 | -35.9646 | -38.4388
22.5003 | -35.9652 | -38.4459
22.5064 | -35.966 | -38.4529
22.5125 | -35.9664 | -38.4599
22.5186 | -35.9669 | -38.4661
22.5247 | -35.9676 | -38.4725
22.5308 | -35.9682 | -38.4797
22.5369 | -35.9687 | -38.4875
22.543 | -35.9693 | -38.4948
22.5491 | -35.97 | -38.5015
22.5552 | -35.9706 | -38.5087
22.5613 | -35.9712 | -38.5154
22.5675 | -35.9719 | -38.5222
22.5736 | -35.9726 | -38.529
22.5797 | -35.9733 | -38.5359
22.5858 | -35.9739 | -38.5429
22.5919 | -35.9746 | -38.5498
22.598 | -35.9752 | -38.5566
22.6041 | -35.9759 | -38.5633
22.6102 | -35.9765 | -38.5704
22.6163 | -35.9772 | -38.5772
22.6224 | -35.9779 | -38.5839
22.6285 | -35.9785 | -38.5906
22.6346 | -35.9791 | -38.5976
22.6407 | -35.9798 | -38.6043
22.6468 | -35.9804 | -38.6115
22.6529 | -35.9811 | -38.6181
22.659 | -35.9818 | -38.6247
22.6651 | -35.9824 | -38.6317
22.6712 | -35.983 | -38.6386
22.6773 | -35.9836 | -38.6457
22.6834 | -35.9843 | -38.6527
22.6895 | -35.985 | -38.6597
22.6956 | -35.9856 | -38.6667
22.7017 | -35.9863 | -38.6737
22.7079 | -35.9869 | -38.6808
22.714 | -35.9875 | -38.688
22.7201 | -35.9881 | -38.6953
22.7262 | -35.9888 | -38.7019
22.7323 | -35.9895 | -38.7089
22.7384 | -35.9902 | -38.7158
22.7445 | -35.9909 | -38.7229
22.7506 | -35.9915 | -38.7301
22.7567 | -35.9922 | -38.7369
22.7628 | -35.9928 | -38.7443
22.7689 | -35.9936 | -38.751
22.775 | -35.9942 | -38.7577
22.7811 | -35.9948 | -38.7642
22.7872 | -35.9955 | -38.7714
22.7933 | -35.9961 | -38.7775
22.7994 | -35.9967 | -38.7843
22.8055 | -35.9974 | -38.7914
22.8116 | -35.9981 | -38.7984
22.8177 | -35.9987 | -38.8053
22.8238 | -35.9993 | -38.8122
22.8299 | -36 | -38.8189
22.836 | -36.0006 | -38.8262
22.8421 | -36.0013 | -38.8329
22.8482 | -36.002 | -38.8398
22.8544 | -36.0027 | -38.8476
22.8605 | -36.0032 | -38.8538
22.8666 | -36.0039 | -38.8608
22.8727 | -36.0044 | -38.8685
22.8788 | -36.0052 | -38.8752
22.8849 | -36.0057 | -38.8825
22.891 | -36.0063 | -38.8891
22.8971 | -36.007 | -38.896
22.9032 | -36.0076 | -38.9026
22.9093 | -36.0083 | -38.9095
22.9154 | -36.0089 | -38.9162
22.9215 | -36.0096 | -38.9232
22.9276 | -36.0102 | -38.9299
22.9337 | -36.0109 | -38.9361
22.9398 | -36.0116 | -38.9431
22.9459 | -36.0122 | -38.9501
22.952 | -36.0128 | -38.9569
22.9581 | -36.0135 | -38.9636
22.9642 | -36.0141 | -38.9704
22.9703 | -36.0147 | -38.977
22.9764 | -36.0154 | -38.9837
22.9825 | -36.016 | -38.9904
22.9886 | -36.0167 | -38.9974
22.9948 | -36.0172 | -39.0045
23.0009 | -36.0179 | -39.0117
23.007 | -36.0185 | -39.0188
23.0131 | -36.0192 | -39.0254
23.0192 | -36.0198 | -39.0326
23.0253 | -36.0205 | -39.0399
23.0314 | -36.0211 | -39.0469
23.0375 | -36.0218 | -39.0539
23.0436 | -36.0224 | -39.0604
23.0497 | -36.023 | -39.0674
23.0558 | -36.0237 | -39.0742
23.0619 | -36.0244 | -39.081
23.068 | -36.025 | -39.088
23.0741 | -36.0257 | -39.0948
23.0802 | -36.0263 | -39.1017
23.0863 | -36.0269 | -39.1086
23.0924 | -36.0275 | -39.1152
23.0985 | -36.0282 | -39.1223
23.1046 | -36.0288 | -39.1292
23.1107 | -36.0295 | -39.1361
23.1168 | -36.0301 | -39.1434
23.1229 | -36.0308 | -39.1504
23.129 | -36.0315 | -39.1571
23.1351 | -36.0321 | -39.1634
23.1413 | -36.0327 | -39.1703
23.1474 | -36.0334 | -39.1769
23.1535 | -36.034 | -39.1839
23.1596 | -36.0347 | -39.191
23.1657 | -36.0354 | -39.1983
23.1718 | -36.0361 | -39.2057
23.1779 | -36.0367 | -39.213
23.184 | -36.0374 | -39.2195
23.1901 | -36.038 | -39.2265
23.1962 | -36.0386 | -39.2333
23.2023 | -36.0392 | -39.24
23.2084 | -36.0398 | -39.2467
23.2145 | -36.0405 | -39.2526
23.2206 | -36.0411 | -39.2595
23.2267 | -36.0418 | -39.2665
23.2328 | -36.0424 | -39.2734
23.2389 | -36.0432 | -39.2804
23.245 | -36.0439 | -39.2871
23.2511 | -36.0445 | -39.2933
23.2572 | -36.0452 | -39.3003
23.2633 | -36.0459 | -39.307
23.2694 | -36.0465 | -39.3141
23.2755 | -36.0472 | -39.3209
23.2817 | -36.0478 | -39.3278
23.2878 | -36.0484 | -39.3342
23.2939 | -36.0491 | -39.3411
23.3 | -36.0497 | -39.3482
23.3061 | -36.0503 | -39.3553
23.3122 | -36.051 | -39.3619
23.3183 | -36.0517 | -39.3686
23.3244 | -36.0524 | -39.3747
23.3305 | -36.053 | -39.3811
23.3366 | -36.0536 | -39.3878
23.3427 | -36.0542 | -39.3944
23.3488 | -36.0549 | -39.4007
23.3549 | -36.0556 | -39.4074
23.361 | -36.0563 | -39.4136
23.3671 | -36.057 | -39.4208
23.3732 | -36.0576 | -39.4272
23.3793 | -36.0582 | -39.4341
23.3854 | -36.0589 | -39.4406
23.3915 | -36.0596 | -39.447
23.3976 | -36.0602 | -39.4537
23.4037 | -36.0609 | -39.461
23.4098 | -36.0616 | -39.4678
23.4159 | -36.0622 | -39.4747
23.422 | -36.0629 | -39.4819
23.4282 | -36.0635 | -39.488
23.4343 | -36.0642 | -39.4948
23.4404 | -36.0649 | -39.5011
23.4465 | -36.0655 | -39.508
23.4526 | -36.0662 | -39.5151
23.4587 | -36.0668 | -39.5215
23.4648 | -36.0674 | -39.5274
23.4709 | -36.0681 | -39.5341
23.477 | -36.0688 | -39.5407
23.4831 | -36.0695 | -39.5475
23.4892 | -36.0701 | -39.5541
23.4953 | -36.0707 | -39.5607
23.5014 | -36.0714 | -39.5671
23.5075 | -36.072 | -39.574
23.5136 | -36.0727 | -39.5814
23.5197 | -36.0734 | -39.5885
23.5258 | -36.074 | -39.595
23.5319 | -36.0747 | -39.6013
23.538 | -36.0753 | -39.6082
23.5441 | -36.076 | -39.6152
23.5502 | -36.0766 | -39.6219
23.5563 | -36.0773 | -39.6285
23.5624 | -36.0779 | -39.6351
23.5686 | -36.0786 | -39.6416
23.5747 | -36.0793 | -39.6482
23.5808 | -36.08 | -39.655
23.5869 | -36.0806 | -39.6614
23.593 | -36.0812 | -39.6683
23.5991 | -36.0819 | -39.675
23.6052 | -36.0827 | -39.6811
23.6113 | -36.0834 | -39.6879
23.6174 | -36.084 | -39.6946
23.6235 | -36.0846 | -39.7013
23.6296 | -36.0852 | -39.708
23.6357 | -36.0859 | -39.7152
23.6418 | -36.0866 | -39.7212
23.6479 | -36.0873 | -39.7275
23.654 | -36.088 | -39.7322
23.6601 | -36.0888 | -39.7398
23.6662 | -36.0897 | -39.7458
23.6723 | -36.0906 | -39.7513
23.6784 | -36.0912 | -39.7578
23.6845 | -36.0919 | -39.7644
23.6906 | -36.0925 | -39.7712
23.6967 | -36.0932 | -39.7777
23.7028 | -36.0938 | -39.784
23.7089 | -36.0945 | -39.7895
23.7151 | -36.0952 | -39.7959
23.7212 | -36.0958 | -39.8027
23.7273 | -36.0964 | -39.8092
23.7334 | -36.0971 | -39.8155
23.7395 | -36.0978 | -39.8226
23.7456 | -36.0984 | -39.8286
23.7517 | -36.0991 | -39.835
23.7578 | -36.0997 | -39.8414
23.7639 | -36.1004 | -39.8481
23.77 | -36.101 | -39.8545
23.7761 | -36.1017 | -39.8613
23.7822 | -36.1024 | -39.8676
23.7883 | -36.103 | -39.8742
23.7944 | -36.1036 | -39.8808
23.8005 | -36.1043 | -39.8873
23.8066 | -36.1049 | -39.8939
23.8127 | -36.1055 | -39.9
23.8188 | -36.1062 | -39.9062
23.8249 | -36.1068 | -39.9127
23.831 | -36.1075 | -39.9192
23.8371 | -36.1081 | -39.9257
23.8432 | -36.1088 | -39.9324
23.8493 | -36.1095 | -39.9387
23.8555 | -36.1101 | -39.9448
23.8616 | -36.1108 | -39.9516
23.8677 | -36.1114 | -39.9583
23.8738 | -36.112 | -39.9652
23.8799 | -36.1127 | -39.9713
23.886 | -36.1134 | -39.9776
23.8921 | -36.1141 | -39.9841
23.8982 | -36.1148 | -39.9909
23.9043 | -36.1154 | -39.9972
23.9104 | -36.1161 | -40.0039
23.9165 | -36.1168 | -40.0101
23.9226 | -36.1174 | -40.0165
23.9287 | -36.1181 | -40.0232
23.9348 | -36.1188 | -40.0297
23.9409 | -36.1195 | -40.0365
23.947 | -36.1202 | -40.0428
23.9531 | -36.1209 | -40.0485
23.9592 | -36.1216 | -40.0549
23.9653 | -36.1222 | -40.0611
23.9714 | -36.1229 | -40.0674
23.9775 | -36.1236 | -40.0738
23.9836 | -36.1243 | -40.0801
23.9897 | -36.125 | -40.0861
23.9958 | -36.1256 | -40.0925
24.002 | -36.1262 | -40.0991
24.0081 | -36.1269 | -40.1056
24.0142 | -36.1276 | -40.112
24.0203 | -36.1283 | -40.1186
24.0264 | -36.1289 | -40.1249
24.0325 | -36.1296 | -40.1312
24.0386 | -36.1303 | -40.1377
24.0447 | -36.131 | -40.1438
24.0508 | -36.1317 | -40.1501
24.0569 | -36.1324 | -40.1562
24.063 | -36.1331 | -40.162
24.0691 | -36.1338 | -40.1681
24.0752 | -36.1345 | -40.1744
24.0813 | -36.1351 | -40.1811
24.0874 | -36.1357 | -40.1876
24.0935 | -36.1363 | -40.1938
24.0996 | -36.137 | -40.1997
24.1057 | -36.1376 | -40.206
24.1118 | -36.1383 | -40.2122
24.1179 | -36.1389 | -40.2187
24.124 | -36.1395 | -40.2252
24.1301 | -36.1402 | -40.2313
24.1362 | -36.1408 | -40.2376
24.1424 | -36.1415 | -40.2441
24.1485 | -36.1422 | -40.2504
24.1546 | -36.1429 | -40.2564
24.1607 | -36.1436 | -40.2625
24.1668 | -36.1444 | -40.2685
24.1729 | -36.145 | -40.275
24.179 | -36.1457 | -40.2814
24.1851 | -36.1463 | -40.2884
24.1912 | -36.147 | -40.2949
24.1973 | -36.1477 | -40.3009
24.2034 | -36.1483 | -40.3071
24.2095 | -36.149 | -40.3136
24.2156 | -36.1496 | -40.3192
24.2217 | -36.1503 | -40.3254
24.2278 | -36.151 | -40.3311
24.2339 | -36.1517 | -40.3371
24.24 | -36.1524 | -40.343
24.2461 | -36.1531 | -40.3497
24.2522 | -36.1537 | -40.3562
24.2583 | -36.1544 | -40.3625
24.2644 | -36.1551 | -40.369
24.2705 | -36.1557 | -40.3748
24.2766 | -36.1564 | -40.3809
24.2827 | -36.157 | -40.3872
24.2889 | -36.1578 | -40.3937
24.295 | -36.1585 | -40.3998
24.3011 | -36.1591 | -40.406
24.3072 | -36.1597 | -40.4115
24.3133 | -36.1604 | -40.418
24.3194 | -36.161 | -40.4249
24.3255 | -36.1617 | -40.4316
24.3316 | -36.1623 | -40.4378
24.3377 | -36.163 | -40.4443
24.3438 | -36.1636 | -40.4499
24.3499 | -36.1643 | -40.4562
24.356 | -36.165 | -40.4626
24.3621 | -36.1656 | -40.4687
24.3682 | -36.1663 | -40.4746
24.3743 | -36.1669 | -40.4808
24.3804 | -36.1676 | -40.4864
24.3865 | -36.1683 | -40.4926
24.3926 | -36.1689 | -40.4989
24.3987 | -36.1696 | -40.5053
24.4048 | -36.1702 | -40.5116
24.4109 | -36.1708 | -40.5173
24.417 | -36.1713 | -40.5231
24.4231 | -36.172 | -40.5292
24.4293 | -36.1727 | -40.5352
24.4354 | -36.1735 | -40.541
24.4415 | -36.174 | -40.5476
24.4476 | -36.1745 | -40.5528
24.4537 | -36.1752 | -40.5587
24.4598 | -36.1758 | -40.565
24.4659 | -36.1765 | -40.5711
24.472 | -36.1772 | -40.5775
24.4781 | -36.1779 | -40.5836
24.4842 | -36.1785 | -40.5893
24.4903 | -36.1791 | -40.596
24.4964 | -36.1797 | -40.6021
24.5025 | -36.1804 | -40.608
24.5086 | -36.181 | -40.614
24.5147 | -36.1816 | -40.6202
24.5208 | -36.1823 | -40.6255
24.5269 | -36.1829 | -40.6318
24.533 | -36.1835 | -40.6379
24.5391 | -36.1842 | -40.6439
24.5452 | -36.1849 | -40.6502
24.5513 | -36.1856 | -40.6562
24.5574 | -36.1863 | -40.6621
24.5635 | -36.1869 | -40.6684
24.5696 | -36.1875 | -40.6744
24.5758 | -36.1883 | -40.6802
24.5819 | -36.1889 | -40.6859
24.588 | -36.1896 | -40.6916
24.5941 | -36.1902 | -40.6979
24.6002 | -36.1909 | -40.704
24.6063 | -36.1916 | -40.7098
24.6124 | -36.1924 | -40.7156
24.6185 | -36.193 | -40.7218
24.6246 | -36.1937 | -40.7272
24.6307 | -36.1944 | -40.7332
24.6368 | -36.195 | -40.7395
24.6429 | -36.1958 | -40.7454
24.649 | -36.1964 | -40.7513
24.6551 | -36.197 | -40.756
24.6612 | -36.1976 | -40.7607
24.6673 | -36.1982 | -40.767
24.6734 | -36.199 | -40.773
24.6795 | -36.1996 | -40.7793
24.6856 | -36.2002 | -40.7855
24.6917 | -36.201 | -40.791
24.6978 | -36.2018 | -40.7961
24.7039 | -36.2024 | -40.8019
24.71 | -36.203 | -40.8079
24.7162 | -36.2037 | -40.8135
24.7223 | -36.2044 | -40.8198
24.7284 | -36.2051 | -40.8248
24.7345 | -36.2057 | -40.8304
24.7406 | -36.2064 | -40.8362
24.7467 | -36.2071 | -40.8425
24.7528 | -36.2078 | -40.8486
24.7589 | -36.2085 | -40.855
24.765 | -36.2091 | -40.8602
24.7711 | -36.2098 | -40.866
24.7772 | -36.2106 | -40.8717
24.7833 | -36.2112 | -40.8774
24.7894 | -36.2119 | -40.8838
24.7955 | -36.2127 | -40.8899
24.8016 | -36.2132 | -40.8951
24.8077 | -36.2139 | -40.9008
24.8138 | -36.2146 | -40.907
24.8199 | -36.2152 | -40.9129
24.826 | -36.2159 | -40.919
24.8321 | -36.2166 | -40.9249
24.8382 | -36.2172 | -40.9306
24.8443 | -36.2178 | -40.9367
24.8504 | -36.2184 | -40.9425
24.8565 | -36.2191 | -40.9483
24.8627 | -36.2198 | -40.9544
24.8688 | -36.2204 | -40.9604
24.8749 | -36.2211 | -40.9662
24.881 | -36.2218 | -40.9723
24.8871 | -36.2225 | -40.9783
24.8932 | -36.2232 | -40.9839
24.8993 | -36.2239 | -40.9895
24.9054 | -36.2246 | -40.9947
24.9115 | -36.2253 | -41.0003
24.9176 | -36.2259 | -41.0061
24.9237 | -36.2266 | -41.0119
24.9298 | -36.2273 | -41.018
24.9359 | -36.228 | -41.024
24.942 | -36.2286 | -41.0289
24.9481 | -36.2293 | -41.035
24.9542 | -36.23 | -41.0409
24.9603 | -36.2307 | -41.0468
24.9664 | -36.2313 | -41.0524
24.9725 | -36.232 | -41.0582
24.9786 | -36.2327 | -41.0633
24.9847 | -36.2333 | -41.0692
24.9908 | -36.2339 | -41.0751
24.9969 | -36.2345 | -41.081
25.0031 | -36.2351 | -41.0869
25.0092 | -36.2357 | -41.0924
25.0153 | -36.2364 | -41.098
25.0214 | -36.2371 | -41.1039
25.0275 | -36.2377 | -41.1097
25.0336 | -36.2384 | -41.1157
25.0397 | -36.2391 | -41.1218
25.0458 | -36.2398 | -41.1267
25.0519 | -36.2405 | -41.1324
25.058 | -36.2411 | -41.1382
25.0641 | -36.2418 | -41.1441
25.0702 | -36.2424 | -41.15
25.0763 | -36.2431 | -41.1557
25.0824 | -36.2437 | -41.1608
25.0885 | -36.2444 | -41.1665
25.0946 | -36.2451 | -41.1726
25.1007 | -36.2457 | -41.1783
25.1068 | -36.2464 | -41.1839
25.1129 | -36.2471 | -41.1898
25.119 | -36.2478 | -41.1952
25.1251 | -36.2485 | -41.2007
25.1312 | -36.2491 | -41.2063
25.1373 | -36.2498 | -41.2118
25.1435 | -36.2505 | -41.2178
25.1496 | -36.2512 | -41.2228
25.1557 | -36.2519 | -41.2284
25.1618 | -36.2526 | -41.2344
25.1679 | -36.2532 | -41.2399
25.174 | -36.2539 | -41.2456
25.1801 | -36.2546 | -41.2516
25.1862 | -36.2552 | -41.2572
25.1923 | -36.2559 | -41.2628
25.1984 | -36.2566 | -41.2687
25.2045 | -36.2572 | -41.2747
25.2106 | -36.2579 | -41.2806
25.2167 | -36.2585 | -41.2868
25.2228 | -36.2591 | -41.292
25.2289 | -36.2598 | -41.2968
25.235 | -36.2604 | -41.303
25.2411 | -36.261 | -41.3082
25.2472 | -36.2618 | -41.3139
25.2533 | -36.2624 | -41.3202
25.2594 | -36.2631 | -41.3254
25.2655 | -36.2638 | -41.3309
25.2716 | -36.2644 | -41.3367
25.2777 | -36.2651 | -41.3424
25.2838 | -36.2657 | -41.3484
25.29 | -36.2664 | -41.3537
25.2961 | -36.267 | -41.3589
25.3022 | -36.2677 | -41.3644
25.3083 | -36.2684 | -41.3696
25.3144 | -36.2691 | -41.3755
25.3205 | -36.2697 | -41.3813
25.3266 | -36.2703 | -41.3865
25.3327 | -36.271 | -41.3921
25.3388 | -36.2717 | -41.3978
25.3449 | -36.2724 | -41.4036
25.351 | -36.273 | -41.4093
25.3571 | -36.2737 | -41.4149
25.3632 | -36.2743 | -41.4199
25.3693 | -36.275 | -41.4255
25.3754 | -36.2756 | -41.4313
25.3815 | -36.2763 | -41.4369
25.3876 | -36.2769 | -41.4426
25.3937 | -36.2775 | -41.4484
25.3998 | -36.2781 | -41.4535
25.4059 | -36.2789 | -41.459
25.412 | -36.2795 | -41.4641
25.4181 | -36.2802 | -41.4698
25.4242 | -36.2809 | -41.4755
25.4304 | -36.2816 | -41.4806
25.4365 | -36.2824 | -41.4858
25.4426 | -36.283 | -41.4916
25.4487 | -36.2836 | -41.497
25.4548 | -36.2842 | -41.5026
25.4609 | -36.2848 | -41.5081
25.467 | -36.2854 | -41.5133
25.4731 | -36.286 | -41.5191
25.4792 | -36.2867 | -41.5251
25.4853 | -36.2874 | -41.5304
25.4914 | -36.288 | -41.5361
25.4975 | -36.2886 | -41.5415
25.5036 | -36.2893 | -41.5469
25.5097 | -36.2899 | -41.552
25.5158 | -36.2906 | -41.5577
25.5219 | -36.2912 | -41.5636
25.528 | -36.2919 | -41.5694
25.5341 | -36.2925 | -41.5748
25.5402 | -36.2931 | -41.5797
25.5463 | -36.2938 | -41.5857
25.5524 | -36.2945 | -41.5913
25.5585 | -36.2951 | -41.5965
25.5646 | -36.2958 | -41.6021
25.5707 | -36.2964 | -41.6079
25.5769 | -36.297 | -41.6128
25.583 | -36.2976 | -41.6184
25.5891 | -36.2982 | -41.6239
25.5952 | -36.2989 | -41.6294
25.6013 | -36.2995 | -41.6351
25.6074 | -36.3001 | -41.6402
25.6135 | -36.3008 | -41.6454
25.6196 | -36.3015 | -41.6509
25.6257 | -36.3022 | -41.6561
25.6318 | -36.3028 | -41.6617
25.6379 | -36.3034 | -41.6677
25.644 | -36.304 | -41.673
25.6501 | -36.3047 | -41.6784
25.6562 | -36.3053 | -41.6837
25.6623 | -36.3059 | -41.689
25.6684 | -36.3066 | -41.6942
25.6745 | -36.3072 | -41.7
25.6806 | -36.308 | -41.7047
25.6867 | -36.3086 | -41.7106
25.6928 | -36.3093 | -41.7168
25.6989 | -36.3099 | -41.7219
25.705 | -36.3105 | -41.7279
25.7111 | -36.3111 | -41.7336
25.7173 | -36.3117 | -41.7392
25.7234 | -36.3124 | -41.7449
25.7295 | -36.3131 | -41.7508
25.7356 | -36.3137 | -41.7569
25.7417 | -36.3144 | -41.7622
25.7478 | -36.315 | -41.7667
25.7539 | -36.3157 | -41.7723
25.76 | -36.3163 | -41.7777
25.7661 | -36.317 | -41.7833
25.7722 | -36.3177 | -41.7888
25.7783 | -36.3183 | -41.7947
25.7844 | -36.3189 | -41.7998
25.7905 | -36.3196 | -41.8052
25.7966 | -36.3202 | -41.8104
25.8027 | -36.3208 | -41.8158
25.8088 | -36.3214 | -41.8214
25.8149 | -36.322 | -41.8266
25.821 | -36.3226 | -41.8313
25.8271 | -36.3233 | -41.8369
25.8332 | -36.3239 | -41.8421
25.8393 | -36.3246 | -41.8479
25.8454 | -36.3252 | -41.8536
25.8515 | -36.3258 | -41.8587
25.8576 | -36.3265 | -41.8636
25.8638 | -36.3271 | -41.8691
25.8699 | -36.3278 | -41.8744
25.876 | -36.3284 | -41.8797
25.8821 | -36.329 | -41.8852
25.8882 | -36.3296 | -41.8898
25.8943 | -36.3302 | -41.8954
25.9004 | -36.3308 | -41.9009
25.9065 | -36.3314 | -41.9063
25.9126 | -36.3321 | -41.9114
25.9187 | -36.3327 | -41.9169
25.9248 | -36.3333 | -41.9223
25.9309 | -36.334 | -41.9276
25.937 | -36.3346 | -41.9334
25.9431 | -36.3353 | -41.9389
25.9492 | -36.3359 | -41.9442
25.9553 | -36.3365 | -41.9495
25.9614 | -36.3371 | -41.9538
25.9675 | -36.3377 | -41.9592
25.9736 | -36.3383 | -41.9649
25.9797 | -36.339 | -41.9703
25.9858 | -36.3396 | -41.9758
25.9919 | -36.3403 | -41.9809
25.998 | -36.3409 | -41.9854
26.0042 | -36.3415 | -41.9905
26.0103 | -36.3422 | -41.9958
26.0164 | -36.3428 | -42.0007
26.0225 | -36.3434 | -42.0067
26.0286 | -36.3441 | -42.012
26.0347 | -36.3447 | -42.0171
26.0408 | -36.3454 | -42.0226
26.0469 | -36.3461 | -42.028
26.053 | -36.3466 | -42.0332
26.0591 | -36.3473 | -42.0386
26.0652 | -36.3478 | -42.0443
26.0713 | -36.3484 | -42.0495
26.0774 | -36.349 | -42.0549
26.0835 | -36.3497 | -42.0604
26.0896 | -36.3504 | -42.0656
26.0957 | -36.3511 | -42.0705
26.1018 | -36.3518 | -42.0763
26.1079 | -36.3524 | -42.0815
26.114 | -36.3531 | -42.0865
26.1201 | -36.3537 | -42.0913
26.1262 | -36.3543 | -42.0966
26.1323 | -36.3549 | -42.1018
26.1384 | -36.3555 | -42.1074
26.1445 | -36.3561 | -42.1126
26.1507 | -36.3567 | -42.1175
26.1568 | -36.3573 | -42.1229
26.1629 | -36.358 | -42.1286
26.169 | -36.3587 | -42.1339
26.1751 | -36.3593 | -42.1392
26.1812 | -36.3598 | -42.1438
26.1873 | -36.3604 | -42.1493
26.1934 | -36.361 | -42.1548
26.1995 | -36.3617 | -42.1601
26.2056 | -36.3623 | -42.1652
26.2117 | -36.3629 | -42.1705
26.2178 | -36.3635 | -42.1752
26.2239 | -36.3641 | -42.1804
26.23 | -36.3648 | -42.1854
26.2361 | -36.3654 | -42.1907
26.2422 | -36.3661 | -42.1953
26.2483 | -36.3667 | -42.2007
26.2544 | -36.3673 | -42.2058
26.2605 | -36.3679 | -42.2109
26.2666 | -36.3685 | -42.2159
26.2727 | -36.3691 | -42.2208
26.2788 | -36.3697 | -42.2254
26.2849 | -36.3704 | -42.2306
26.2911 | -36.3711 | -42.2353
26.2972 | -36.3717 | -42.2405
26.3033 | -36.3724 | -42.2457
26.3094 | -36.3731 | -42.251
26.3155 | -36.3737 | -42.2563
26.3216 | -36.3744 | -42.262
26.3277 | -36.375 | -42.2675
26.3338 | -36.3756 | -42.2725
26.3399 | -36.3761 | -42.2783
26.346 | -36.3767 | -42.2833
26.3521 | -36.3773 | -42.2889
26.3582 | -36.378 | -42.2941
26.3643 | -36.3786 | -42.2992
26.3704 | -36.3792 | -42.3042
26.3765 | -36.3798 | -42.3097
26.3826 | -36.3803 | -42.3141
26.3887 | -36.381 | -42.3192
26.3948 | -36.3816 | -42.3245
26.4009 | -36.3822 | -42.3306
26.407 | -36.3828 | -42.3356
26.4131 | -36.3835 | -42.3413
26.4192 | -36.384 | -42.3462
26.4253 | -36.3847 | -42.3509
26.4314 | -36.3853 | -42.3568
26.4376 | -36.386 | -42.3618
26.4437 | -36.3866 | -42.3672
26.4498 | -36.3871 | -42.3728
26.4559 | -36.3877 | -42.3774
26.462 | -36.3884 | -42.3829
26.4681 | -36.389 | -42.3882
26.4742 | -36.3895 | -42.3933
26.4803 | -36.3901 | -42.3982
26.4864 | -36.3908 | -42.403
26.4925 | -36.3915 | -42.408
26.4986 | -36.3921 | -42.4132
26.5047 | -36.3927 | -42.4185
26.5108 | -36.3933 | -42.4236
26.5169 | -36.3939 | -42.4292
26.523 | -36.3945 | -42.434
26.5291 | -36.3951 | -42.4387
26.5352 | -36.3956 | -42.4437
26.5413 | -36.3962 | -42.4489
26.5474 | -36.3967 | -42.4545
26.5535 | -36.3974 | -42.4597
26.5596 | -36.3979 | -42.4656
26.5657 | -36.3986 | -42.4705
26.5718 | -36.3992 | -42.475
26.578 | -36.3999 | -42.4796
26.5841 | -36.4005 | -42.4848
26.5902 | -36.4011 | -42.4905
26.5963 | -36.4017 | -42.4958
26.6024 | -36.4023 | -42.5014
26.6085 | -36.4029 | -42.5067
26.6146 | -36.4035 | -42.5119
26.6207 | -36.4041 | -42.5169
26.6268 | -36.4048 | -42.5222
26.6329 | -36.4053 | -42.5269
26.639 | -36.4059 | -42.532
26.6451 | -36.4065 | -42.5374
26.6512 | -36.4071 | -42.5428
26.6573 | -36.4077 | -42.5483
26.6634 | -36.4083 | -42.5527
26.6695 | -36.4089 | -42.5582
26.6756 | -36.4096 | -42.5634
26.6817 | -36.4101 | -42.5697
26.6878 | -36.4108 | -42.5751
26.6939 | -36.4114 | -42.58
26.7 | -36.4119 | -42.5852
26.7061 | -36.4126 | -42.5904
26.7122 | -36.4133 | -42.5954
26.7183 | -36.4139 | -42.6009
26.7245 | -36.4145 | -42.6063
26.7306 | -36.4151 | -42.6113
26.7367 | -36.4156 | -42.6171
26.7428 | -36.4162 | -42.6229
26.7489 | -36.4167 | -42.6283
26.755 | -36.4173 | -42.6333
26.7611 | -36.4178 | -42.6389
26.7672 | -36.4184 | -42.6438
26.7733 | -36.419 | -42.6485
26.7794 | -36.4196 | -42.6536
26.7855 | -36.4202 | -42.6587
26.7916 | -36.4208 | -42.6636
26.7977 | -36.4215 | -42.6689
26.8038 | -36.4221 | -42.6739
26.8099 | -36.4228 | -42.6786
26.816 | -36.4235 | -42.6843
26.8221 | -36.4242 | -42.6889
26.8282 | -36.4247 | -42.6937
26.8343 | -36.4253 | -42.6985
26.8404 | -36.4258 | -42.7036
26.8465 | -36.4265 | -42.7088
26.8526 | -36.427 | -42.7143
26.8587 | -36.4277 | -42.7192
26.8649 | -36.4282 | -42.7239
26.871 | -36.4288 | -42.7294
26.8771 | -36.4294 | -42.7344
26.8832 | -36.4299 | -42.7396
26.8893 | -36.4305 | -42.7453
26.8954 | -36.4309 | -42.7504
26.9015 | -36.4316 | -42.7561
26.9076 | -36.4323 | -42.7606
26.9137 | -36.4328 | -42.7659
26.9198 | -36.4335 | -42.7716
26.9259 | -36.4341 | -42.7775
26.932 | -36.4346 | -42.7827
26.9381 | -36.4351 | -42.7877
26.9442 | -36.4355 | -42.7923
26.9503 | -36.4362 | -42.7976
26.9564 | -36.4367 | -42.8028
26.9625 | -36.4373 | -42.8081
26.9686 | -36.4378 | -42.8138
26.9747 | -36.4383 | -42.8189
26.9808 | -36.4388 | -42.8246
26.9869 | -36.4393 | -42.8299
26.993 | -36.4399 | -42.8354
26.9991 | -36.4405 | -42.8407
27.0052 | -36.4411 | -42.8466
27.0114 | -36.4416 | -42.8517
27.0175 | -36.4421 | -42.8572
27.0236 | -36.4427 | -42.8627
27.0297 | -36.4433 | -42.8674
27.0358 | -36.4438 | -42.8727
27.0419 | -36.4444 | -42.8778
27.048 | -36.445 | -42.8827
27.0541 | -36.4454 | -42.8877
27.0602 | -36.446 | -42.8929
27.0663 | -36.4466 | -42.8981
27.0724 | -36.4471 | -42.9036
27.0785 | -36.4476 | -42.9095
27.0846 | -36.448 | -42.9137
27.0907 | -36.4486 | -42.9184
27.0968 | -36.449 | -42.9239
27.1029 | -36.4496 | -42.9294
27.109 | -36.45 | -42.9351
27.1151 | -36.4506 | -42.9403
27.1212 | -36.4511 | -42.9457
27.1273 | -36.4517 | -42.9511
27.1334 | -36.4523 | -42.9554
27.1395 | -36.4528 | -42.9605
27.1456 | -36.4535 | -42.9662
27.1518 | -36.4541 | -42.9716
27.1579 | -36.4546 | -42.9769
27.164 | -36.4551 | -42.9822
27.1701 | -36.4556 | -42.9874
27.1762 | -36.4562 | -42.9921
27.1823 | -36.4568 | -42.9965
27.1884 | -36.4573 | -43.0023
27.1945 | -36.4578 | -43.0077
27.2006 | -36.4584 | -43.0131
27.2067 | -36.459 | -43.0192
27.2128 | -36.4596 | -43.0245
27.2189 | -36.4602 | -43.0297
27.225 | -36.4607 | -43.035
27.2311 | -36.4612 | -43.04
27.2372 | -36.4617 | -43.0458
27.2433 | -36.4623 | -43.0509
27.2494 | -36.4629 | -43.057
27.2555 | -36.4634 | -43.0623
27.2616 | -36.4639 | -43.0675
27.2677 | -36.4645 | -43.0734
27.2738 | -36.4652 | -43.0789
27.2799 | -36.4657 | -43.0839
27.286 | -36.4663 | -43.089
27.2921 | -36.4668 | -43.0941
27.2983 | -36.4673 | -43.0982
27.3044 | -36.4679 | -43.1035
27.3105 | -36.4684 | -43.1086
27.3166 | -36.469 | -43.1135
27.3227 | -36.4695 | -43.119
27.3288 | -36.4702 | -43.1247
27.3349 | -36.4706 | -43.1303
27.341 | -36.4711 | -43.1353
27.3471 | -36.4717 | -43.1401
27.3532 | -36.4722 | -43.145
27.3593 | -36.4728 | -43.1505
27.3654 | -36.4734 | -43.1558
27.3715 | -36.4738 | -43.161
27.3776 | -36.4743 | -43.1664
27.3837 | -36.4749 | -43.1713
27.3898 | -36.4754 | -43.1754
27.3959 | -36.476 | -43.1801
27.402 | -36.4765 | -43.1846
27.4081 | -36.477 | -43.1897
27.4142 | -36.4775 | -43.1952
27.4203 | -36.4781 | -43.2006
27.4264 | -36.4786 | -43.206
27.4325 | -36.4791 | -43.2118
27.4387 | -36.4795 | -43.2159
27.4448 | -36.4801 | -43.2215
27.4509 | -36.4807 | -43.2263
27.457 | -36.4813 | -43.2315
27.4631 | -36.4818 | -43.2369
27.4692 | -36.4823 | -43.2423
27.4753 | -36.4826 | -43.2473
27.4814 | -36.4831 | -43.2527
27.4875 | -36.4837 | -43.2574
27.4936 | -36.4842 | -43.2627
27.4997 | -36.4846 | -43.2671
27.5058 | -36.4851 | -43.2728
27.5119 | -36.4856 | -43.278
27.518 | -36.4861 | -43.2826
27.5241 | -36.4865 | -43.2877
27.5302 | -36.487 | -43.2927
27.5363 | -36.4876 | -43.2982
27.5424 | -36.488 | -43.3025
27.5485 | -36.4885 | -43.3076
27.5546 | -36.489 | -43.312
27.5607 | -36.4895 | -43.3165
27.5668 | -36.4901 | -43.3216
27.5729 | -36.4906 | -43.3267
27.5791 | -36.491 | -43.3329
27.5852 | -36.4913 | -43.3373
27.5913 | -36.492 | -43.3425
27.5974 | -36.4927 | -43.3465
27.6035 | -36.4932 | -43.3538
27.6096 | -36.494 | -43.3583
27.6157 | -36.4945 | -43.364
27.6218 | -36.4951 | -43.369
27.6279 | -36.4955 | -43.3757
27.634 | -36.496 | -43.3799
27.6401 | -36.4965 | -43.3853
27.6462 | -36.4971 | -43.3899
27.6523 | -36.4976 | -43.3942
27.6584 | -36.4981 | -43.3991
27.6645 | -36.4987 | -43.4041
27.6706 | -36.4992 | -43.41
27.6767 | -36.4996 | -43.4144
27.6828 | -36.5 | -43.4193
27.6889 | -36.5005 | -43.4247
27.695 | -36.5009 | -43.4308
27.7011 | -36.5013 | -43.4363
27.7072 | -36.5018 | -43.4414
27.7133 | -36.5023 | -43.4465
27.7194 | -36.5028 | -43.452
27.7256 | -36.5034 | -43.4577
27.7317 | -36.5038 | -43.4633
27.7378 | -36.5043 | -43.4682
27.7439 | -36.5048 | -43.4742
27.75 | -36.5054 | -43.4785
27.7561 | -36.5058 | -43.4842
27.7622 | -36.5062 | -43.4895
27.7683 | -36.5068 | -43.495
27.7744 | -36.5073 | -43.4999
27.7805 | -36.5079 | -43.5053
27.7866 | -36.5084 | -43.51
27.7927 | -36.5086 | -43.515
27.7988 | -36.5091 | -43.5205
27.8049 | -36.5097 | -43.5252
27.811 | -36.5101 | -43.5308
27.8171 | -36.5106 | -43.536
27.8232 | -36.5111 | -43.5409
27.8293 | -36.5116 | -43.5446
27.8354 | -36.5121 | -43.5493
27.8415 | -36.5127 | -43.5546
27.8476 | -36.5132 | -43.5603
27.8537 | -36.5136 | -43.5657
27.8598 | -36.514 | -43.5706
27.866 | -36.5145 | -43.576
27.8721 | -36.515 | -43.5808
27.8782 | -36.5156 | -43.5853
27.8843 | -36.5161 | -43.5908
27.8904 | -36.5166 | -43.5954
27.8965 | -36.517 | -43.6004
27.9026 | -36.5175 | -43.606
27.9087 | -36.518 | -43.6115
27.9148 | -36.5185 | -43.6159
27.9209 | -36.519 | -43.6213
27.927 | -36.5195 | -43.6267
27.9331 | -36.5199 | -43.6325
27.9392 | -36.5203 | -43.6374
27.9453 | -36.5208 | -43.6437
27.9514 | -36.5212 | -43.65
27.9575 | -36.5217 | -43.6549
27.9636 | -36.5222 | -43.6602
27.9697 | -36.5226 | -43.6656
27.9758 | -36.523 | -43.671
27.9819 | -36.5235 | -43.6766
27.988 | -36.524 | -43.6812
27.9941 | -36.5244 | -43.6865
28.0002 | -36.5248 | -43.6914
28.0063 | -36.5253 | -43.6962
28.0125 | -36.5257 | -43.7016
28.0186 | -36.5261 | -43.7064
28.0247 | -36.5267 | -43.711
28.0308 | -36.5272 | -43.7174
28.0369 | -36.5276 | -43.7225
28.043 | -36.5281 | -43.7291
28.0491 | -36.5284 | -43.7335
28.0552 | -36.5286 | -43.7385
28.0613 | -36.5292 | -43.7437
28.0674 | -36.5296 | -43.7489
28.0735 | -36.5301 | -43.755

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -37.2872 | -49.1851
21.4908 | -36.1313 | -41.3682
16.4948 | -35.3565 | -33.3698
12.6602 | -34.8627 | -26.369
9.7171 | -34.55 | -20.7316
7.4582 | -34.3548 | -16.0118
5.7244 | -34.2038 | -12.1831
4.3936 | -34.0959 | -9.0996
3.3722 | -34.068 | -6.6251
2.5883 | -34.0531 | -4.7549
1.9866 | -34.0322 | -3.264
1.5248 | -34.0315 | -2.0712
1.1703 | -34.0274 | -1.2648
1 | -34.0149 | -0.8571
0.89822 | -34.0229 | -0.53227
0.68941 | -34.0118 | 0.13851
0.52919 | -33.9863 | 0.6749
0.40612 | -33.9917 | 1.3474
0.31171 | -33.9689 | 1.8489
0.23928 | -33.9677 | 2.5911
0.18358 | -33.9818 | 3.0442
0.14091 | -33.9691 | 3.9484
0.10819 | -33.985 | 5.1202
0.083017 | -34.0007 | 6.6155
0.063705 | -34.0515 | 8.4846
0.048923 | -34.1017 | 10.9503
0.037479 | -34.1981 | 14.1352
0.028801 | -34.3662 | 18.2206
0.022078 | -34.6802 | 23.69
0.017023 | -35.1757 | 30.0763
0.013018 | -35.9846 | 37.7401
0.010014 | -37.0992 | 45.752

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003719 | -10.5897
0.007439 | -34.5632
0.011158 | -38.5092
0.014877 | -42.6224
0.018596 | -44.0223
0.022316 | -45.2269
0.026035 | -46.2509
0.029754 | -47.5951
0.033473 | -49.9467
0.037193 | -49.8808
0.040912 | -50.5091
0.044631 | -52.1795
0.048351 | -54.0968
0.05207 | -54.8014
0.055789 | -55.6371
0.059508 | -57.6392
0.063228 | -57.0156
0.066947 | -57.3119
0.070666 | -58.5092
0.074385 | -57.9322
0.078105 | -58.3939
0.081824 | -59.5934
0.085543 | -60.7888
0.089262 | -61.2253
0.092982 | -61.2412
0.096701 | -61.4419
0.10042 | -61.8224
0.10414 | -61.7977
0.10786 | -62.5933
0.11158 | -62.4114
0.1153 | -62.6431
0.11902 | -63.1514
0.12274 | -64.0333
0.12646 | -64.1799
0.13017 | -64.6522
0.13389 | -64.9652
0.13761 | -64.9818
0.14133 | -64.8064
0.14505 | -65.3401
0.14877 | -67.0386
0.15249 | -67.2854
0.15621 | -67.0829
0.15993 | -67.2539
0.16365 | -68.6224
0.16737 | -68.1819
0.17109 | -68.8223
0.17481 | -68.8021
0.17852 | -68.7926
0.18224 | -68.3585
0.18596 | -68.2268
0.18968 | -68.4832
0.1934 | -68.6478
0.19712 | -68.9483
0.20084 | -68.5176
0.20456 | -69.3174
0.20828 | -69.6884
0.212 | -70.2267
0.21572 | -70.9061
0.21944 | -70.0872
0.22316 | -70.545
0.22687 | -71.1122
0.23059 | -70.795
0.23431 | -72.2284
0.23803 | -71.3757
0.24175 | -71.8241
0.24547 | -72.7887
0.24919 | -73.3433
0.25291 | -74.2178
0.25663 | -72.5837
0.26035 | -72.7523
0.26407 | -73.7239
0.26779 | -73.9999
0.27151 | -73.7739
0.27523 | -73.0115
0.27894 | -73.2352
0.28266 | -74.6239
0.28638 | -75.1097
0.2901 | -74.0767
0.29382 | -74.1921
0.29754 | -74.4143
0.30126 | -75.0612
0.30498 | -75.628
0.3087 | -75.2747
0.31242 | -75.1233
0.31614 | -75.7541
0.31986 | -75.8008
0.32358 | -75.2858
0.3273 | -74.9259
0.33102 | -76.1522
0.33473 | -76.5778
0.33845 | -76.0795
0.34217 | -75.2844
0.34589 | -75.0672
0.34961 | -75.8306
0.35333 | -76.6286
0.35705 | -75.2597
0.36077 | -75.1477
0.36449 | -76.012
0.36821 | -75.8593
0.37193 | -75.4619
0.37565 | -74.863
0.37937 | -75.3449
0.38309 | -75.7271
0.3868 | -76.3856
0.39052 | -77.2432
0.39424 | -76.5116
0.39796 | -77.1747
0.40168 | -77.738
0.4054 | -77.3458
0.40912 | -77.8489
0.41284 | -78.2556
0.41656 | -78.059
0.42028 | -78.0277
0.424 | -77.5319
0.42772 | -77.8667
0.43144 | -77.0365
0.43516 | -77.1035
0.43887 | -78.6554
0.44259 | -77.9243
0.44631 | -78.403
0.45003 | -77.7115
0.45375 | -77.1369
0.45747 | -77.4174
0.46119 | -77.8821
0.46491 | -78.6062
0.46863 | -79.4007
0.47235 | -78.0225
0.47607 | -78.4281
0.47979 | -80.195
0.48351 | -80.1797
0.48722 | -78.7503
0.49094 | -78.3831
0.49466 | -78.9948
0.49838 | -78.9293
0.5021 | -78.9555
0.50582 | -79.2115
0.50954 | -79.9179
0.51326 | -78.6246
0.51698 | -79.1044
0.5207 | -80.1883
0.52442 | -79.369
0.52814 | -79.5034
0.53186 | -79.8293
0.53558 | -80.4102
0.53929 | -80.4409
0.54301 | -80.3135
0.54673 | -80.3419
0.55045 | -80.456
0.55417 | -80.1965
0.55789 | -79.9326
0.56161 | -79.59
0.56533 | -79.1793
0.56905 | -79.298
0.57277 | -80.0021
0.57649 | -80.807
0.58021 | -81.1514
0.58393 | -81.1167
0.58764 | -81.0412
0.59136 | -81.2729
0.59508 | -81.1642
0.5988 | -80.5291
0.60252 | -80.2428
0.60624 | -80.6601
0.60996 | -81.3909
0.61368 | -81.3713
0.6174 | -80.7612
0.62112 | -80.3096
0.62484 | -80.7373
0.62856 | -81.5204
0.63228 | -81.2801
0.63599 | -80.5974
0.63971 | -80.3283
0.64343 | -80.5664
0.64715 | -81.0916
0.65087 | -81.7507
0.65459 | -82.0458
0.65831 | -82.0957
0.66203 | -82.272
0.66575 | -82.5425
0.66947 | -83.0344
0.67319 | -83.0843
0.67691 | -82.5466
0.68063 | -82.5942
0.68435 | -83.076
0.68807 | -83.0712
0.69178 | -82.8491
0.6955 | -82.9898
0.69922 | -83.1774
0.70294 | -83.092
0.70666 | -82.768
0.71038 | -82.689
0.7141 | -82.7503
0.71782 | -82.6175
0.72154 | -82.7277
0.72526 | -83.1622
0.72898 | -83.3135
0.7327 | -83.4325
0.73642 | -83.7141
0.74013 | -83.5651
0.74385 | -83.1706
0.74757 | -82.9668
0.75129 | -83.3563
0.75501 | -83.9818
0.75873 | -84.0285
0.76245 | -83.6363
0.76617 | -83.1666
0.76989 | -83.0187
0.77361 | -82.9551
0.77733 | -82.7525
0.78105 | -82.5792
0.78477 | -82.7766
0.78848 | -83.2388
0.7922 | -83.3303
0.79592 | -83.2319
0.79964 | -83.1054
0.80336 | -83.1331
0.80708 | -83.5244
0.8108 | -83.9801
0.81452 | -84.1834
0.81824 | -84.2174
0.82196 | -84.7461
0.82568 | -85.3144
0.8294 | -85.0011
0.83312 | -84.3527
0.83684 | -83.9358
0.84056 | -83.9555
0.84427 | -84.4453
0.84799 | -84.7056
0.85171 | -84.8617
0.85543 | -85.1315
0.85915 | -85.0467
0.86287 | -84.4463
0.86659 | -84.3979
0.87031 | -84.9589
0.87403 | -85.0653
0.87775 | -84.7079
0.88147 | -84.4389
0.88519 | -84.2689
0.8889 | -84.139
0.89263 | -84.2318
0.89634 | -84.4504
0.90006 | -84.792
0.90378 | -85.2834
0.9075 | -85.7894
0.91122 | -86.2793
0.91494 | -86.4757
0.91866 | -86.1726
0.92238 | -85.4406
0.9261 | -84.8059
0.92982 | -84.9477
0.93354 | -85.7546
0.93726 | -86.3718
0.94098 | -86.0466
0.94469 | -85.5737
0.94841 | -85.578
0.95213 | -85.7668
0.95585 | -85.3923
0.95957 | -84.6569
0.96329 | -84.5735
0.96701 | -84.8614
0.97073 | -85.2582
0.97445 | -85.9374
0.97817 | -86.5518
0.98189 | -86.4205
0.98561 | -86.0185
0.98933 | -85.4065
0.99304 | -83.0237
0.99676 | -79.9168
1.0005 | -78.8743
1.0042 | -80.9848
1.0079 | -84.4094
1.0116 | -85.9172
1.0154 | -86.2718
1.0191 | -86.5679
1.0228 | -86.7331
1.0265 | -86.5448
1.0302 | -86.4763
1.034 | -86.4296
1.0377 | -86.1533
1.0414 | -86.0899
1.0451 | -86.592
1.0488 | -86.7105
1.0526 | -86.0655
1.0563 | -85.8377
1.06 | -86.417
1.0637 | -87.1343
1.0674 | -87.2819
1.0712 | -87.0089
1.0749 | -86.9819
1.0786 | -87.1839
1.0823 | -87.3117
1.086 | -86.962
1.0897 | -86.3608
1.0935 | -85.9567
1.0972 | -85.7379
1.1009 | -85.5177
1.1046 | -85.4845
1.1083 | -85.8763
1.1121 | -86.5483
1.1158 | -87.0423
1.1195 | -87.1818
1.1232 | -87.26
1.1269 | -87.3522
1.1307 | -87.2271
1.1344 | -86.9237
1.1381 | -86.7255
1.1418 | -86.8203
1.1455 | -87.0997
1.1493 | -87.2385
1.153 | -87.2007
1.1567 | -87.2132
1.1604 | -87.3108
1.1641 | -87.4515
1.1679 | -87.6898
1.1716 | -88.0002
1.1753 | -88.2599
1.179 | -88.4214
1.1827 | -88.4413
1.1864 | -88.2202
1.1902 | -87.8397
1.1939 | -87.5539
1.1976 | -87.5443
1.2013 | -87.7652
1.205 | -88.0728
1.2088 | -88.3845
1.2125 | -88.5141
1.2162 | -88.3603
1.2199 | -88.2387
1.2236 | -88.3729
1.2274 | -88.5314
1.2311 | -88.347
1.2348 | -87.8701
1.2385 | -87.4903
1.2422 | -87.4467
1.246 | -87.7568
1.2497 | -88.2987
1.2534 | -88.7479
1.2571 | -88.6708
1.2608 | -87.9966
1.2646 | -87.1175
1.2683 | -86.4741
1.272 | -86.1804
1.2757 | -86.0747
1.2794 | -86.0825
1.2831 | -86.2435
1.2869 | -86.5249
1.2906 | -86.8907
1.2943 | -87.3099
1.298 | -87.7128
1.3017 | -87.9562
1.3055 | -87.9638
1.3092 | -87.8562
1.3129 | -87.832
1.3166 | -88.0048
1.3203 | -88.3448
1.3241 | -88.8086
1.3278 | -89.3402
1.3315 | -89.6167
1.3352 | -89.472
1.3389 | -89.2545
1.3427 | -89.2451
1.3464 | -89.3107
1.3501 | -89.2335
1.3538 | -88.9794
1.3575 | -88.6828
1.3613 | -88.5135
1.365 | -88.5859
1.3687 | -88.8407
1.3724 | -89.0468
1.3761 | -89.1617
1.3798 | -89.3783
1.3836 | -89.7375
1.3873 | -90.0118
1.391 | -89.8333
1.3947 | -89.2612
1.3984 | -88.78
1.4022 | -88.5711
1.4059 | -88.4768
1.4096 | -88.4106
1.4133 | -88.5004
1.417 | -88.716
1.4208 | -88.9563
1.4245 | -89.3455
1.4282 | -89.8436
1.4319 | -90.0678
1.4356 | -89.8332
1.4394 | -89.3899
1.4431 | -89.1139
1.4468 | -89.0806
1.4505 | -89.0375
1.4542 | -88.9422
1.458 | -88.9809
1.4617 | -89.1827
1.4654 | -89.5317
1.4691 | -89.933
1.4728 | -90.0687
1.4766 | -89.8619
1.4803 | -89.578
1.484 | -89.4454
1.4877 | -89.497
1.4914 | -89.6647
1.4951 | -89.7668
1.4989 | -89.7143
1.5026 | -89.6849
1.5063 | -89.7835
1.51 | -89.9104
1.5137 | -89.9468
1.5175 | -89.8754
1.5212 | -89.8366
1.5249 | -89.8554
1.5286 | -89.8974
1.5323 | -90.0505
1.5361 | -90.3033
1.5398 | -90.5415
1.5435 | -90.6493
1.5472 | -90.5273
1.5509 | -90.25
1.5547 | -89.9899
1.5584 | -89.8295
1.5621 | -89.7538
1.5658 | -89.7567
1.5695 | -89.886
1.5733 | -90.14
1.577 | -90.4762
1.5807 | -90.7782
1.5844 | -90.8892
1.5881 | -90.8101
1.5918 | -90.6734
1.5956 | -90.5108
1.5993 | -90.323
1.603 | -90.1426
1.6067 | -90.1458
1.6104 | -90.197
1.6142 | -90.2855
1.6179 | -90.3136
1.6216 | -90.2426
1.6253 | -90.1597
1.629 | -90.1519
1.6328 | -90.2094
1.6365 | -90.2911
1.6402 | -90.4316
1.6439 | -90.6827
1.6476 | -90.969
1.6514 | -91.1256
1.6551 | -91.0741
1.6588 | -90.8378
1.6625 | -90.5284
1.6662 | -90.3026
1.67 | -90.2726
1.6737 | -90.4412
1.6774 | -90.7171
1.6811 | -90.9558
1.6848 | -91.0165
1.6885 | -90.8804
1.6923 | -90.6982
1.696 | -90.5893
1.6997 | -90.563
1.7034 | -90.55
1.7071 | -90.5087
1.7109 | -90.4905
1.7146 | -90.5615
1.7183 | -90.7333
1.722 | -90.9204
1.7257 | -91.0234
1.7295 | -91.0542
1.7332 | -91.0599
1.7369 | -91.0372
1.7406 | -90.9844
1.7443 | -90.9419
1.7481 | -90.9288
1.7518 | -90.9251
1.7555 | -90.9061
1.7592 | -90.8744
1.7629 | -90.8124
1.7667 | -90.7331
1.7704 | -90.7125
1.7741 | -90.8214
1.7778 | -91.0325
1.7815 | -91.1913
1.7852 | -91.1341
1.789 | -90.8945
1.7927 | -90.6461
1.7964 | -90.5469
1.8001 | -90.6396
1.8038 | -90.8351
1.8076 | -91.0194
1.8113 | -91.1344
1.815 | -91.21
1.8187 | -91.2565
1.8224 | -91.2335
1.8262 | -91.1457
1.8299 | -91.0668
1.8336 | -91.0684
1.8373 | -91.1717
1.841 | -91.3298
1.8448 | -91.4993
1.8485 | -91.6256
1.8522 | -91.6556
1.8559 | -91.6065
1.8596 | -91.5372
1.8634 | -91.5221
1.8671 | -91.5893
1.8708 | -91.6931
1.8745 | -91.7278
1.8782 | -91.6733
1.882 | -91.611
1.8857 | -91.6337
1.8894 | -91.712
1.8931 | -91.72
1.8968 | -91.5135
1.9005 | -91.0689
1.9043 | -90.5204
1.908 | -90.0155
1.9117 | -89.6235
1.9154 | -89.3742
1.9191 | -89.2682
1.9229 | -89.3159
1.9266 | -89.5403
1.9303 | -89.9405
1.934 | -90.4456
1.9377 | -90.9257
1.9415 | -91.3264
1.9452 | -91.6493
1.9489 | -91.8966
1.9526 | -92.0127
1.9563 | -91.9625
1.9601 | -91.8079
1.9638 | -91.6806
1.9675 | -91.7051
1.9712 | -91.899
1.9749 | -92.1412
1.9787 | -92.2486
1.9824 | -91.992
1.9861 | -91.1853
1.9898 | -89.9171
1.9935 | -88.521
1.9972 | -87.5634
2.001 | -87.389
2.0047 | -88.0643
2.0084 | -89.3857
2.0121 | -90.9134
2.0158 | -92.1603
2.0196 | -92.9021
2.0233 | -93.1565
2.027 | -93.2028
2.0307 | -93.0943
2.0344 | -92.901
2.0382 | -92.6812
2.0419 | -92.4534
2.0456 | -92.1327
2.0493 | -91.7136
2.053 | -91.329
2.0568 | -91.156
2.0605 | -91.2901
2.0642 | -91.685
2.0679 | -92.1498
2.0716 | -92.4741
2.0754 | -92.6265
2.0791 | -92.7481
2.0828 | -92.9158
2.0865 | -93.1277
2.0902 | -93.3643
2.0939 | -93.5565
2.0977 | -93.5722
2.1014 | -93.3832
2.1051 | -93.0791
2.1088 | -92.7657
2.1125 | -92.5334
2.1163 | -92.4419
2.12 | -92.4847
2.1237 | -92.5604
2.1274 | -92.6272
2.1311 | -92.7736
2.1349 | -93.0502
2.1386 | -93.327
2.1423 | -93.4258
2.146 | -93.1455
2.1497 | -92.9307
2.1535 | -92.7105
2.1572 | -92.5507
2.1609 | -92.4961
2.1646 | -92.5323
2.1683 | -92.6174
2.1721 | -92.7052
2.1758 | -92.7687
2.1795 | -92.7829
2.1832 | -92.7432
2.1869 | -92.6624
2.1907 | -92.5318
2.1944 | -92.3631
2.1981 | -92.197
2.2018 | -92.0715
2.2055 | -92.0026
2.2092 | -91.9905
2.213 | -92.0422
2.2167 | -92.1687
2.2204 | -92.3568
2.2241 | -92.5617
2.2278 | -92.7178
2.2316 | -92.7964
2.2353 | -92.8229
2.239 | -92.8678
2.2427 | -92.9976
2.2464 | -93.2137
2.2502 | -93.4638
2.2539 | -93.6507
2.2576 | -93.721
2.2613 | -93.6433
2.265 | -93.4184
2.2688 | -93.079
2.2725 | -92.7
2.2762 | -92.3607
2.2799 | -92.1419
2.2836 | -92.0883
2.2874 | -92.213
2.2911 | -92.4542
2.2948 | -92.7045
2.2985 | -92.8598
2.3022 | -92.9087
2.3059 | -92.8831
2.3097 | -92.8198
2.3134 | -92.7404
2.3171 | -92.6635
2.3208 | -92.6244
2.3245 | -92.6571
2.3283 | -92.7655
2.332 | -92.9483
2.3357 | -93.1967
2.3394 | -93.4589
2.3431 | -93.6664
2.3469 | -93.7693
2.3506 | -93.7717
2.3543 | -93.7372
2.358 | -93.7132
2.3617 | -93.7001
2.3655 | -93.695
2.3692 | -93.6987
2.3729 | -93.7209
2.3766 | -93.7374
2.3803 | -93.7454
2.3841 | -93.7543
2.3878 | -93.777
2.3915 | -93.8149
2.3952 | -93.835
2.3989 | -93.8326
2.4026 | -93.8154
2.4064 | -93.7789
2.4101 | -93.7041
2.4138 | -93.5991
2.4175 | -93.5062
2.4212 | -93.4489
2.425 | -93.4553
2.4287 | -93.5113
2.4324 | -93.575
2.4361 | -93.6426
2.4398 | -93.7201
2.4436 | -93.7904
2.4473 | -93.8233
2.451 | -93.8083
2.4547 | -93.7727
2.4584 | -93.7678
2.4622 | -93.7974
2.4659 | -93.7908
2.4696 | -93.7021
2.4733 | -93.5576
2.477 | -93.4202
2.4808 | -93.3469
2.4845 | -93.3603
2.4882 | -93.4128
2.4919 | -93.4642
2.4956 | -93.5073
2.4993 | -93.5665
2.5031 | -93.6426
2.5068 | -93.7086
2.5105 | -93.7335
2.5142 | -93.7071
2.5179 | -93.6772
2.5217 | -93.6752
2.5254 | -93.6698
2.5291 | -93.608
2.5328 | -93.4179
2.5365 | -93.1134
2.5403 | -92.767
2.544 | -92.4725
2.5477 | -92.2864
2.5514 | -92.1827
2.5551 | -92.1089
2.5589 | -92.0224
2.5626 | -91.9267
2.5663 | -91.8825
2.57 | -91.9196
2.5737 | -92.0175
2.5775 | -92.1406
2.5812 | -92.2892
2.5849 | -92.4772
2.5886 | -92.6797
2.5923 | -92.8759
2.596 | -93.048
2.5998 | -93.2069
2.6035 | -93.3832
2.6072 | -93.5452
2.6109 | -93.6643
2.6146 | -93.7465
2.6184 | -93.7739
2.6221 | -93.7307
2.6258 | -93.6242
2.6295 | -93.4996
2.6332 | -93.3953
2.637 | -93.3137
2.6407 | -93.2428
2.6444 | -93.2015
2.6481 | -93.2425
2.6518 | -93.3906
2.6556 | -93.6032
2.6593 | -93.797
2.663 | -93.8804
2.6667 | -93.8186
2.6704 | -93.6517
2.6742 | -93.4695
2.6779 | -93.3614
2.6816 | -93.4981
2.6853 | -93.6126
2.689 | -93.7848
2.6928 | -93.9617
2.6965 | -94.0923
2.7002 | -94.1536
2.7039 | -94.1676
2.7076 | -94.1681
2.7113 | -94.1773
2.7151 | -94.2173
2.7188 | -94.2839
2.7225 | -94.3846
2.7262 | -94.5103
2.7299 | -94.6306
2.7337 | -94.7098
2.7374 | -94.7353
2.7411 | -94.6901
2.7448 | -94.5809
2.7485 | -94.4418
2.7523 | -94.316
2.756 | -94.2468
2.7597 | -94.2274
2.7634 | -94.2139
2.7671 | -94.1805
2.7709 | -94.1219
2.7746 | -94.0425
2.7783 | -93.9449
2.782 | -93.8381
2.7857 | -93.7202
2.7895 | -93.6436
2.7932 | -93.6599
2.7969 | -93.7795
2.8006 | -93.9877
2.8043 | -94.2258
2.808 | -94.4376
2.8118 | -94.5769
2.8155 | -94.6182
2.8192 | -94.5569
2.8229 | -94.4409
2.8266 | -94.3327
2.8304 | -94.286
2.8341 | -94.3353
2.8378 | -94.4643
2.8415 | -94.6173
2.8452 | -94.7187
2.849 | -94.7403
2.8527 | -94.6936
2.8564 | -94.6196
2.8601 | -94.57
2.8638 | -94.5589
2.8676 | -94.5779
2.8713 | -94.6137
2.875 | -94.6589
2.8787 | -94.6929
2.8824 | -94.7071
2.8862 | -94.7046
2.8899 | -94.7095
2.8936 | -94.7167
2.8973 | -94.7084
2.901 | -94.6804
2.9047 | -94.6429
2.9085 | -94.6099
2.9122 | -94.618
2.9159 | -94.6811
2.9196 | -94.7796
2.9233 | -94.897
2.9271 | -95.0077
2.9308 | -95.0938
2.9345 | -95.1453
2.9382 | -95.1617
2.9419 | -95.1626
2.9457 | -95.1302
2.9494 | -95.0707
2.9531 | -94.9967
2.9568 | -94.9395
2.9605 | -94.9128
2.9643 | -94.9117
2.968 | -94.9166
2.9717 | -94.9022
2.9754 | -94.8711
2.9791 | -94.8424
2.9829 | -94.8234
2.9866 | -94.7745
2.9903 | -94.671
2.994 | -94.519
2.9977 | -94.3641
3.0015 | -94.2691
3.0052 | -94.2725
3.0089 | -94.3918
3.0126 | -94.6093
3.0163 | -94.8838
3.02 | -95.1621
3.0238 | -95.3901
3.0275 | -95.5245
3.0312 | -95.5474
3.0349 | -95.4894
3.0386 | -95.3876
3.0424 | -95.2963
3.0461 | -95.2563
3.0498 | -95.2876
3.0535 | -95.3832
3.0572 | -95.4952
3.061 | -95.5763
3.0647 | -95.5857
3.0684 | -95.5194
3.0721 | -95.4073
3.0758 | -95.2757
3.0796 | -95.1684
3.0833 | -95.1014
3.087 | -95.0671
3.0907 | -95.0589
3.0944 | -95.0691
3.0982 | -95.0904
3.1019 | -95.1219
3.1056 | -95.1723
3.1093 | -95.2229
3.113 | -95.2767
3.1167 | -95.3166
3.1205 | -95.3595
3.1242 | -95.4176
3.1279 | -95.4714
3.1316 | -95.5008
3.1353 | -95.5092
3.1391 | -95.5175
3.1428 | -95.5604
3.1465 | -95.6421
3.1502 | -95.7633
3.1539 | -95.8996
3.1577 | -96.0036
3.1614 | -96.0318
3.1651 | -95.9697
3.1688 | -95.8026
3.1725 | -95.5553
3.1763 | -95.2738
3.18 | -95.0037
3.1837 | -94.7828
3.1874 | -94.6406
3.1911 | -94.5767
3.1949 | -94.5763
3.1986 | -94.6231
3.2023 | -94.7172
3.206 | -94.8568
3.2097 | -95.0106
3.2134 | -95.1647
3.2172 | -95.3094
3.2209 | -95.4542
3.2246 | -95.5947
3.2283 | -95.713
3.232 | -95.8171
3.2358 | -95.8783
3.2395 | -95.8916
3.2432 | -95.8797
3.2469 | -95.8642
3.2506 | -95.851
3.2544 | -95.8384
3.2581 | -95.837
3.2618 | -95.8317
3.2655 | -95.8096
3.2692 | -95.7537
3.273 | -95.6667
3.2767 | -95.5716
3.2804 | -95.4803
3.2841 | -95.4154
3.2878 | -95.3818
3.2916 | -95.3972
3.2953 | -95.4735
3.299 | -95.599
3.3027 | -95.7692
3.3064 | -95.9697
3.3102 | -96.1821
3.3139 | -96.3637
3.3176 | -96.4977
3.3213 | -96.558
3.325 | -96.55
3.3287 | -96.484
3.3325 | -96.3794
3.3362 | -96.2758
3.3399 | -96.1976
3.3436 | -96.1394
3.3473 | -96.1051
3.3511 | -96.0692
3.3548 | -96.0279
3.3585 | -95.9633
3.3622 | -95.8849
3.3659 | -95.7953
3.3697 | -95.7075
3.3734 | -95.6208
3.3771 | -95.559
3.3808 | -95.5198
3.3845 | -95.5091
3.3883 | -95.5275
3.392 | -95.5923
3.3957 | -95.6887
3.3994 | -95.8115
3.4031 | -95.9313
3.4069 | -96.0199
3.4106 | -96.0625
3.4143 | -96.0602
3.418 | -96.0162
3.4217 | -95.9687
3.4254 | -95.942
3.4292 | -95.96
3.4329 | -96.0115
3.4366 | -96.0963
3.4403 | -96.1877
3.444 | -96.2505
3.4478 | -96.257
3.4515 | -96.2128
3.4552 | -96.1277
3.4589 | -96.0423
3.4626 | -95.976
3.4664 | -95.9504
3.4701 | -95.9765
3.4738 | -96.0404
3.4775 | -96.1228
3.4812 | -96.2078
3.485 | -96.2785
3.4887 | -96.3373
3.4924 | -96.374
3.4961 | -96.3758
3.4998 | -96.3473
3.5036 | -96.3093
3.5073 | -96.2715
3.511 | -96.2351
3.5147 | -96.1774
3.5184 | -96.1092
3.5221 | -96.0376
3.5259 | -95.9824
3.5296 | -95.9554
3.5333 | -95.952
3.537 | -95.9616
3.5407 | -95.9773
3.5445 | -96.0001
3.5482 | -96.0406
3.5519 | -96.0978
3.5556 | -96.1523
3.5593 | -96.1712
3.5631 | -96.1741
3.5668 | -96.168
3.5705 | -96.1633
3.5742 | -96.1628
3.5779 | -96.1588
3.5817 | -96.1483
3.5854 | -96.1424
3.5891 | -96.14
3.5928 | -96.1438
3.5965 | -96.1495
3.6003 | -96.1536
3.604 | -96.1494
3.6077 | -96.1417
3.6114 | -96.133
3.6151 | -96.1287
3.6188 | -96.1214
3.6226 | -96.1205
3.6263 | -96.1387
3.63 | -96.166
3.6337 | -96.2056
3.6374 | -96.2487
3.6412 | -96.3007
3.6449 | -96.3603
3.6486 | -96.4071
3.6523 | -96.4456
3.656 | -96.4756
3.6598 | -96.5133
3.6635 | -96.5611
3.6672 | -96.6309
3.6709 | -96.715
3.6746 | -96.8022
3.6784 | -96.8809
3.6821 | -96.9256
3.6858 | -96.9419
3.6895 | -96.9332
3.6932 | -96.9242
3.697 | -96.9131
3.7007 | -96.9118
3.7044 | -96.9189
3.7081 | -96.9084
3.7118 | -96.8774
3.7155 | -96.8215
3.7193 | -96.746
3.723 | -96.6414
3.7267 | -96.5208
3.7304 | -96.4024
3.7341 | -96.3
3.7379 | -96.2245
3.7416 | -96.1939
3.7453 | -96.2048
3.749 | -96.2519
3.7527 | -96.2914
3.7565 | -96.3103
3.7602 | -96.3102
3.7639 | -96.2832
3.7676 | -96.236
3.7713 | -96.1866
3.7751 | -96.1575
3.7788 | -96.1697
3.7825 | -96.2413
3.7862 | -96.3611
3.7899 | -96.5084
3.7937 | -96.6655
3.7974 | -96.8019
3.8011 | -96.8908
3.8048 | -96.9293
3.8085 | -96.9103
3.8123 | -96.8539
3.816 | -96.7852
3.8197 | -96.7073
3.8234 | -96.6504
3.8271 | -96.6319
3.8308 | -96.6515
3.8346 | -96.6981
3.8383 | -96.7564
3.842 | -96.815
3.8457 | -96.865
3.8494 | -96.9032
3.8532 | -96.9325
3.8569 | -96.9659
3.8606 | -97.0113
3.8643 | -97.0614
3.868 | -97.1112
3.8718 | -97.1464
3.8755 | -97.1643
3.8792 | -97.1488
3.8829 | -97.0997
3.8866 | -97.0064
3.8904 | -96.9023
3.8941 | -96.7975
3.8978 | -96.7124
3.9015 | -96.6642
3.9052 | -96.6485
3.909 | -96.6598
3.9127 | -96.703
3.9164 | -96.7628
3.9201 | -96.8177
3.9238 | -96.8626
3.9275 | -96.8895
3.9313 | -96.8932
3.935 | -96.8924
3.9387 | -96.881
3.9424 | -96.8646
3.9461 | -96.8445
3.9499 | -96.8392
3.9536 | -96.8292
3.9573 | -96.8237
3.961 | -96.8287
3.9647 | -96.8304
3.9685 | -96.7366
3.9722 | -96.5743
3.9759 | -96.3614
3.9796 | -96.0525
3.9833 | -95.6721
3.9871 | -95.262
3.9908 | -94.8966
3.9945 | -94.6366
3.9982 | -94.52
4.0019 | -94.5709
4.0057 | -94.7899
4.0094 | -95.1534
4.0131 | -95.5965
4.0168 | -96.0603
4.0205 | -96.4809
4.0243 | -96.7998
4.028 | -96.9969
4.0317 | -97.1285
4.0354 | -97.1311
4.0391 | -97.0427
4.0428 | -96.974
4.0466 | -96.9392
4.0503 | -96.9562
4.054 | -97.018
4.0577 | -97.1231
4.0614 | -97.2546
4.0652 | -97.3986
4.0689 | -97.5384
4.0726 | -97.6522
4.0763 | -97.7441
4.08 | -97.7982
4.0838 | -97.8238
4.0875 | -97.8249
4.0912 | -97.8034
4.0949 | -97.758
4.0986 | -97.6974
4.1024 | -97.6069
4.1061 | -97.4957
4.1098 | -97.3784
4.1135 | -97.2668
4.1172 | -97.1815
4.121 | -97.1355
4.1247 | -97.1302
4.1284 | -97.1592
4.1321 | -97.2021
4.1358 | -97.2415
4.1395 | -97.2703
4.1433 | -97.2977
4.147 | -97.3371
4.1507 | -97.396
4.1544 | -97.4838
4.1581 | -97.5913
4.1619 | -97.6978
4.1656 | -97.7777
4.1693 | -97.7974
4.173 | -97.756
4.1767 | -97.6668
4.1805 | -97.5263
4.1842 | -97.357
4.1879 | -97.1827
4.1916 | -97.0126
4.1953 | -96.8804
4.1991 | -96.7999
4.2028 | -96.7785
4.2065 | -96.8141
4.2102 | -96.891
4.2139 | -97.0054
4.2177 | -97.1188
4.2214 | -97.2085
4.2251 | -97.2757
4.2288 | -97.3144
4.2325 | -97.3408
4.2362 | -97.3508
4.24 | -97.3556
4.2437 | -97.3695
4.2474 | -97.3882
4.2511 | -97.3972
4.2548 | -97.3952
4.2586 | -97.3686
4.2623 | -97.3233
4.266 | -97.2658
4.2697 | -97.1976
4.2734 | -97.1376
4.2772 | -97.085
4.2809 | -97.0532
4.2846 | -97.0381
4.2883 | -97.0395
4.292 | -97.066
4.2958 | -97.088
4.2995 | -97.1102
4.3032 | -97.134
4.3069 | -97.157
4.3106 | -97.183
4.3144 | -97.2075
4.3181 | -97.2396
4.3218 | -97.2796
4.3255 | -97.3342
4.3292 | -97.3959
4.3329 | -97.456
4.3367 | -97.5148
4.3404 | -97.5624
4.3441 | -97.6116
4.3478 | -97.6645
4.3515 | -97.7226
4.3553 | -97.7848
4.359 | -97.8509
4.3627 | -97.9115
4.3664 | -97.9499
4.3701 | -97.9766
4.3739 | -97.9858
4.3776 | -97.9846
4.3813 | -97.9699
4.385 | -97.945
4.3887 | -97.9163
4.3925 | -97.8969
4.3962 | -97.8698
4.3999 | -97.8479
4.4036 | -97.8448
4.4073 | -97.8488
4.4111 | -97.8579
4.4148 | -97.8603
4.4185 | -97.8529
4.4222 | -97.8392
4.4259 | -97.8196
4.4296 | -97.8057
4.4334 | -97.7959
4.4371 | -97.7917
4.4408 | -97.7859
4.4445 | -97.7759
4.4482 | -97.7524
4.452 | -97.7324
4.4557 | -97.7169
4.4594 | -97.6917
4.4631 | -97.6592
4.4668 | -97.6326
4.4706 | -97.606
4.4743 | -97.5706
4.478 | -97.5336
4.4817 | -97.4944
4.4854 | -97.4552
4.4892 | -97.4225
4.4929 | -97.4051
4.4966 | -97.4002
4.5003 | -97.422
4.504 | -97.4817
4.5078 | -97.5683
4.5115 | -97.6689
4.5152 | -97.7827
4.5189 | -97.898
4.5226 | -97.9936
4.5264 | -98.0733
4.5301 | -98.1368
4.5338 | -98.175
4.5375 | -98.1971
4.5412 | -98.2079
4.5449 | -98.2174
4.5487 | -98.2391
4.5524 | -98.2808
4.5561 | -98.3435
4.5598 | -98.4205
4.5635 | -98.4939
4.5673 | -98.5575
4.571 | -98.591
4.5747 | -98.58
4.5784 | -98.5402
4.5821 | -98.4759
4.5859 | -98.3816
4.5896 | -98.2772
4.5933 | -98.1701
4.597 | -98.0643
4.6007 | -97.9713
4.6045 | -97.8919
4.6082 | -97.8321
4.6119 | -97.7977
4.6156 | -97.7879
4.6193 | -97.7996
4.6231 | -97.8275
4.6268 | -97.8705
4.6305 | -97.9223
4.6342 | -97.9823
4.6379 | -98.0446
4.6416 | -98.1003
4.6454 | -98.1495
4.6491 | -98.1798
4.6528 | -98.1827
4.6565 | -98.1503
4.6602 | -98.0743
4.664 | -97.9583
4.6677 | -97.8223
4.6714 | -97.69
4.6751 | -97.5822
4.6788 | -97.5105
4.6826 | -97.4852
4.6863 | -97.5105
4.69 | -97.5811
4.6937 | -97.6797
4.6974 | -97.7828
4.7012 | -97.8778
4.7049 | -97.9519
4.7086 | -97.997
4.7123 | -98.0148
4.716 | -98.0159
4.7198 | -98.0109
4.7235 | -98.0146
4.7272 | -98.0371
4.7309 | -98.0875
4.7346 | -98.1629
4.7383 | -98.2502
4.7421 | -98.3357
4.7458 | -98.4191
4.7495 | -98.4947
4.7532 | -98.5544
4.7569 | -98.5947
4.7607 | -98.6177
4.7644 | -98.634
4.7681 | -98.6478
4.7718 | -98.6566
4.7755 | -98.6593
4.7793 | -98.6666
4.783 | -98.658
4.7867 | -98.642
4.7904 | -98.6136
4.7941 | -98.5794
4.7979 | -98.5435
4.8016 | -98.5028
4.8053 | -98.475
4.809 | -98.457
4.8127 | -98.4482
4.8165 | -98.442
4.8202 | -98.4407
4.8239 | -98.4417
4.8276 | -98.4461
4.8313 | -98.4444
4.835 | -98.4371
4.8388 | -98.4234
4.8425 | -98.4065
4.8462 | -98.3811
4.8499 | -98.3529
4.8536 | -98.3195
4.8574 | -98.302
4.8611 | -98.284
4.8648 | -98.2803
4.8685 | -98.2869
4.8722 | -98.3084
4.876 | -98.3409
4.8797 | -98.3732
4.8834 | -98.3977
4.8871 | -98.4144
4.8908 | -98.4108
4.8946 | -98.3707
4.8983 | -98.3051
4.902 | -98.2141
4.9057 | -98.1075
4.9094 | -98.0018
4.9132 | -97.9072
4.9169 | -97.8393
4.9206 | -97.7986
4.9243 | -97.7906
4.928 | -97.811
4.9318 | -97.8604
4.9355 | -97.9227
4.9392 | -97.9938
4.9429 | -98.0556
4.9466 | -98.1044
4.9503 | -98.1445
4.9541 | -98.1717
4.9578 | -98.173
4.9615 | -98.059
4.9652 | -97.8935
4.9689 | -97.7162
4.9727 | -97.4952
4.9764 | -97.2258
4.9801 | -96.9281
4.9838 | -96.6277
4.9875 | -96.3522
4.9913 | -96.1284
4.995 | -95.973
4.9987 | -95.9062
5.0024 | -95.9404
5.0061 | -96.0689
5.0099 | -96.2795
5.0136 | -96.5546
5.0173 | -96.8692
5.021 | -97.1976
5.0247 | -97.5087
5.0285 | -97.7654
5.0322 | -97.9689
5.0359 | -98.127
5.0396 | -98.2983
5.0433 | -98.3902
5.047 | -98.4065
5.0508 | -98.4278
5.0545 | -98.4574
5.0582 | -98.4924
5.0619 | -98.5267
5.0656 | -98.5613
5.0694 | -98.6002
5.0731 | -98.6537
5.0768 | -98.7201
5.0805 | -98.7832
5.0842 | -98.8415
5.088 | -98.8938
5.0917 | -98.9247
5.0954 | -98.9336
5.0991 | -98.9054
5.1028 | -98.8528
5.1066 | -98.7863
5.1103 | -98.7083
5.114 | -98.6326
5.1177 | -98.5727
5.1214 | -98.532
5.1252 | -98.5171
5.1289 | -98.524
5.1326 | -98.5516
5.1363 | -98.5915
5.14 | -98.6374
5.1437 | -98.6816
5.1475 | -98.7099
5.1512 | -98.7218
5.1549 | -98.7293
5.1586 | -98.7387
5.1623 | -98.7562
5.1661 | -98.7769
5.1698 | -98.8013
5.1735 | -98.8191
5.1772 | -98.8412
5.1809 | -98.8561
5.1847 | -98.8638
5.1884 | -98.8565
5.1921 | -98.8433
5.1958 | -98.819
5.1995 | -98.7889
5.2033 | -98.7624
5.207 | -98.7483
5.2107 | -98.752
5.2144 | -98.7698
5.2181 | -98.8012
5.2219 | -98.8468
5.2256 | -98.9016
5.2293 | -98.9635
5.233 | -99.024
5.2367 | -99.0741
5.2405 | -99.1101
5.2442 | -99.1299
5.2479 | -99.134
5.2516 | -99.1168
5.2553 | -99.0752
5.259 | -99.0216
5.2628 | -98.9672
5.2665 | -98.9187
5.2702 | -98.8785
5.2739 | -98.8671
5.2776 | -98.8897
5.2814 | -98.9364
5.2851 | -99.0041
5.2888 | -99.0942
5.2925 | -99.2067
5.2962 | -99.3131
5.3 | -99.4147
5.3037 | -99.4987
5.3074 | -99.5509
5.3111 | -99.5677
5.3148 | -99.5587
5.3186 | -99.5306
5.3223 | -99.4841
5.326 | -99.4269
5.3297 | -99.3741
5.3334 | -99.3243
5.3372 | -99.2769
5.3409 | -99.2437
5.3446 | -99.2329
5.3483 | -99.2402
5.352 | -99.2596
5.3557 | -99.2891
5.3595 | -99.3334
5.3632 | -99.3901
5.3669 | -99.4417
5.3706 | -99.4858
5.3743 | -99.5163
5.3781 | -99.5379
5.3818 | -99.5405
5.3855 | -99.5269
5.3892 | -99.4972
5.3929 | -99.4524
5.3967 | -99.3957
5.4004 | -99.3292
5.4041 | -99.2606
5.4078 | -99.1959
5.4115 | -99.1419
5.4153 | -99.1047
5.419 | -99.0924
5.4227 | -99.0947
5.4264 | -99.1051
5.4301 | -99.1214
5.4339 | -99.1462
5.4376 | -99.1695
5.4413 | -99.1798
5.445 | -99.1785
5.4487 | -99.1733
5.4524 | -99.1669
5.4562 | -99.1619
5.4599 | -99.1577
5.4636 | -99.1507
5.4673 | -99.1342
5.471 | -99.1203
5.4748 | -99.107
5.4785 | -99.0966
5.4822 | -99.0839
5.4859 | -99.0767
5.4896 | -99.0785
5.4934 | -99.0808
5.4971 | -99.0828
5.5008 | -99.0837
5.5045 | -99.0734
5.5082 | -99.0624
5.512 | -99.0559
5.5157 | -99.0438
5.5194 | -99.0249
5.5231 | -99.01
5.5268 | -98.9987
5.5306 | -98.9902
5.5343 | -98.9759
5.538 | -98.9684
5.5417 | -98.973
5.5454 | -98.9825
5.5492 | -98.9968
5.5529 | -99.0203
5.5566 | -99.0463
5.5603 | -99.0785
5.564 | -99.1124
5.5677 | -99.1486
5.5715 | -99.1856
5.5752 | -99.2289
5.5789 | -99.2802
5.5826 | -99.3276
5.5863 | -99.3746
5.5901 | -99.4242
5.5938 | -99.4649
5.5975 | -99.4924
5.6012 | -99.5146
5.6049 | -99.5238
5.6087 | -99.5158
5.6124 | -99.499
5.6161 | -99.477
5.6198 | -99.4545
5.6235 | -99.4315
5.6273 | -99.4038
5.631 | -99.3798
5.6347 | -99.3656
5.6384 | -99.3553
5.6421 | -99.3485
5.6459 | -99.3468
5.6496 | -99.3487
5.6533 | -99.3525
5.657 | -99.356
5.6607 | -99.3529
5.6644 | -99.3543
5.6682 | -99.3641
5.6719 | -99.3727
5.6756 | -99.3795
5.6793 | -99.3747
5.683 | -99.3547
5.6868 | -99.3274
5.6905 | -99.2885
5.6942 | -99.2437
5.6979 | -99.1904
5.7016 | -99.1269
5.7054 | -99.0533
5.7091 | -98.9822
5.7128 | -98.9132
5.7165 | -98.8409
5.7202 | -98.7632
5.724 | -98.688
5.7277 | -98.62
5.7314 | -98.5584
5.7351 | -98.5035
5.7388 | -98.4536
5.7426 | -98.413
5.7463 | -98.3916
5.75 | -98.3851
5.7537 | -98.3988
5.7574 | -98.4414
5.7611 | -98.5124
5.7649 | -98.6121
5.7686 | -98.7212
5.7723 | -98.8202
5.776 | -98.9096
5.7797 | -98.9789
5.7835 | -99.0252
5.7872 | -99.0484
5.7909 | -99.0465
5.7946 | -99.0238
5.7983 | -98.9904
5.8021 | -98.9626
5.8058 | -98.9339
5.8095 | -98.9085
5.8132 | -98.8932
5.8169 | -98.8832
5.8207 | -98.8893
5.8244 | -98.9078
5.8281 | -98.9289
5.8318 | -98.9516
5.8355 | -98.9762
5.8393 | -99.0016
5.843 | -99.0246
5.8467 | -99.0414
5.8504 | -99.0614
5.8541 | -99.0914
5.8579 | -99.1197
5.8616 | -99.1415
5.8653 | -99.16
5.869 | -99.1811
5.8727 | -99.2076
5.8764 | -99.2364
5.8802 | -99.2738
5.8839 | -99.32
5.8876 | -99.3565
5.8913 | -99.3926
5.895 | -99.4308
5.8988 | -99.4504
5.9025 | -99.4742
5.9062 | -99.4904
5.9099 | -99.4968
5.9136 | -99.4978
5.9174 | -99.4936
5.9211 | -99.4855
5.9248 | -99.4768
5.9285 | -99.4644
5.9322 | -99.4526
5.936 | -99.4491
5.9397 | -99.4499
5.9434 | -99.4587
5.9471 | -99.4768
5.9508 | -99.4973
5.9546 | -99.5177
5.9583 | -99.5388
5.962 | -99.5601
5.9657 | -99.5783
5.9694 | -99.5711
5.9731 | -99.5506
5.9769 | -99.5125
5.9806 | -99.4537
5.9843 | -99.3778
5.988 | -99.2931
5.9917 | -99.2023
5.9955 | -99.117
5.9992 | -99.0495
6.0029 | -99.0007
6.0066 | -98.9726
6.0103 | -98.9666
6.0141 | -98.9752
6.0178 | -99.0004
6.0215 | -99.0403
6.0252 | -99.0868
6.0289 | -99.1373
6.0327 | -99.1931
6.0364 | -99.25
6.0401 | -99.3104
6.0438 | -99.3731
6.0475 | -99.4358
6.0513 | -99.4898
6.055 | -99.5379
6.0587 | -99.5869
6.0624 | -99.633
6.0661 | -99.675
6.0698 | -99.7216
6.0736 | -99.7676
6.0773 | -99.8013
6.081 | -99.84
6.0847 | -99.8796
6.0884 | -99.9178
6.0922 | -99.9466
6.0959 | -99.9713
6.0996 | -99.9919
6.1033 | -100.0021
6.107 | -100.0087
6.1108 | -100.0024
6.1145 | -99.9851
6.1182 | -99.9586
6.1219 | -99.9167
6.1256 | -99.8657
6.1294 | -99.8155
6.1331 | -99.7615
6.1368 | -99.7099
6.1405 | -99.6672
6.1442 | -99.629
6.148 | -99.5967
6.1517 | -99.5746
6.1554 | -99.5638
6.1591 | -99.5572
6.1628 | -99.5611
6.1665 | -99.5727
6.1703 | -99.5992
6.174 | -99.63
6.1777 | -99.6625
6.1814 | -99.6994
6.1851 | -99.737
6.1889 | -99.7672
6.1926 | -99.7947
6.1963 | -99.823
6.2 | -99.8393
6.2037 | -99.8393
6.2075 | -99.83
6.2112 | -99.8252
6.2149 | -99.8213
6.2186 | -99.8108
6.2223 | -99.7917
6.2261 | -99.7686
6.2298 | -99.7458
6.2335 | -99.7164
6.2372 | -99.6848
6.2409 | -99.6557
6.2447 | -99.6244
6.2484 | -99.5977
6.2521 | -99.5742
6.2558 | -99.5536
6.2595 | -99.5446
6.2633 | -99.5453
6.267 | -99.5548
6.2707 | -99.5778
6.2744 | -99.611
6.2781 | -99.6454
6.2818 | -99.6913
6.2856 | -99.7426
6.2893 | -99.7949
6.293 | -99.8481
6.2967 | -99.9003
6.3004 | -99.9389
6.3042 | -99.9488
6.3079 | -99.9492
6.3116 | -99.9414
6.3153 | -99.9295
6.319 | -99.9141
6.3228 | -99.9004
6.3265 | -99.8917
6.3302 | -99.8924
6.3339 | -99.894
6.3376 | -99.9038
6.3414 | -99.9238
6.3451 | -99.9544
6.3488 | -99.9821
6.3525 | -100.0076
6.3562 | -100.0308
6.36 | -100.0497
6.3637 | -100.0525
6.3674 | -100.0282
6.3711 | -99.9921
6.3748 | -99.9335
6.3785 | -99.8476
6.3823 | -99.7468
6.386 | -99.6438
6.3897 | -99.5397
6.3934 | -99.4407
6.3971 | -99.3542
6.4009 | -99.2721
6.4046 | -99.206
6.4083 | -99.1584
6.412 | -99.1329
6.4157 | -99.1198
6.4195 | -99.1185
6.4232 | -99.1325
6.4269 | -99.1598
6.4306 | -99.193
6.4343 | -99.2243
6.4381 | -99.2552
6.4418 | -99.2856
6.4455 | -99.319
6.4492 | -99.3493
6.4529 | -99.3841
6.4567 | -99.427
6.4604 | -99.4775
6.4641 | -99.5408
6.4678 | -99.6072
6.4715 | -99.6731
6.4752 | -99.7471
6.479 | -99.8199
6.4827 | -99.8867
6.4864 | -99.9507
6.4901 | -100.0105
6.4938 | -100.0627
6.4976 | -100.1125
6.5013 | -100.1489
6.505 | -100.1695
6.5087 | -100.1821
6.5124 | -100.1959
6.5162 | -100.1994
6.5199 | -100.1874
6.5236 | -100.1673
6.5273 | -100.1411
6.531 | -100.1162
6.5348 | -100.0852
6.5385 | -100.0511
6.5422 | -100.0183
6.5459 | -99.9886
6.5496 | -99.9657
6.5534 | -99.9405
6.5571 | -99.9165
6.5608 | -99.8983
6.5645 | -99.8803
6.5682 | -99.8661
6.5719 | -99.8585
6.5757 | -99.8528
6.5794 | -99.8505
6.5831 | -99.8516
6.5868 | -99.854
6.5905 | -99.8559
6.5943 | -99.8673
6.598 | -99.8891
6.6017 | -99.9174
6.6054 | -99.9494
6.6091 | -99.9762
6.6129 | -99.9975
6.6166 | -100.0234
6.6203 | -100.0553
6.624 | -100.0807
6.6277 | -100.1037
6.6315 | -100.1248
6.6352 | -100.147
6.6389 | -100.1638
6.6426 | -100.1778
6.6463 | -100.1907
6.6501 | -100.195
6.6538 | -100.1929
6.6575 | -100.1863
6.6612 | -100.1762
6.6649 | -100.1599
6.6687 | -100.1381
6.6724 | -100.1084
6.6761 | -100.0806
6.6798 | -100.0552
6.6835 | -100.0368
6.6872 | -100.0233
6.691 | -100.0185
6.6947 | -100.0158
6.6984 | -100.0167
6.7021 | -100.0308
6.7058 | -100.0535
6.7096 | -100.0735
6.7133 | -100.0958
6.717 | -100.1173
6.7207 | -100.1316
6.7244 | -100.1415
6.7282 | -100.1492
6.7319 | -100.1544
6.7356 | -100.1526
6.7393 | -100.1454
6.743 | -100.137
6.7468 | -100.1223
6.7505 | -100.1081
6.7542 | -100.0966
6.7579 | -100.0879
6.7616 | -100.0788
6.7654 | -100.0776
6.7691 | -100.0817
6.7728 | -100.0945
6.7765 | -100.1057
6.7802 | -100.118
6.7839 | -100.1298
6.7877 | -100.1374
6.7914 | -100.1372
6.7951 | -100.1398
6.7988 | -100.1416
6.8025 | -100.143
6.8063 | -100.1469
6.81 | -100.1569
6.8137 | -100.1594
6.8174 | -100.169
6.8211 | -100.1801
6.8249 | -100.1981
6.8286 | -100.2087
6.8323 | -100.2176
6.836 | -100.2206
6.8397 | -100.2148
6.8435 | -100.2016
6.8472 | -100.1868
6.8509 | -100.1726
6.8546 | -100.1608
6.8583 | -100.1514
6.8621 | -100.143
6.8658 | -100.1332
6.8695 | -100.1317
6.8732 | -100.1292
6.8769 | -100.1268
6.8806 | -100.1265
6.8844 | -100.131
6.8881 | -100.1406
6.8918 | -100.1498
6.8955 | -100.1556
6.8992 | -100.1608
6.903 | -100.1689
6.9067 | -100.1807
6.9104 | -100.1948
6.9141 | -100.2102
6.9178 | -100.2157
6.9216 | -100.2229
6.9253 | -100.2292
6.929 | -100.238
6.9327 | -100.2433
6.9364 | -100.2504
6.9402 | -100.2538
6.9439 | -100.2494
6.9476 | -100.2385
6.9513 | -100.2144
6.955 | -100.1895
6.9588 | -100.1676
6.9625 | -100.1424
6.9662 | -100.1105
6.9699 | -100.0655
6.9736 | -100.0012
6.9774 | -99.9464
6.9811 | -99.8924
6.9848 | -99.8372
6.9885 | -99.7774
6.9922 | -99.7249
6.9959 | -99.6912
6.9997 | -99.6597
7.0034 | -99.6399
7.0071 | -99.6342
7.0108 | -99.6364
7.0145 | -99.648
7.0183 | -99.6625
7.022 | -99.6759
7.0257 | -99.6863
7.0294 | -99.6943
7.0331 | -99.7026
7.0369 | -99.7093
7.0406 | -99.7193
7.0443 | -99.7378
7.048 | -99.766
7.0517 | -99.8094
7.0555 | -99.8705
7.0592 | -99.9375
7.0629 | -100.0131
7.0666 | -100.1007
7.0703 | -100.1934
7.0741 | -100.2836
7.0778 | -100.3643
7.0815 | -100.4421
7.0852 | -100.5105
7.0889 | -100.5597
7.0926 | -100.5823
7.0964 | -100.5928
7.1001 | -100.588
7.1038 | -100.5613
7.1075 | -100.5327
7.1112 | -100.4955
7.115 | -100.454
7.1187 | -100.4171
7.1224 | -100.3836
7.1261 | -100.3615
7.1298 | -100.3509
7.1336 | -100.3544
7.1373 | -100.3708
7.141 | -100.3962
7.1447 | -100.4316
7.1484 | -100.468
7.1522 | -100.5088
7.1559 | -100.5456
7.1596 | -100.5783
7.1633 | -100.6022
7.167 | -100.6132
7.1708 | -100.6175
7.1745 | -100.6134
7.1782 | -100.5961
7.1819 | -100.5685
7.1856 | -100.5406
7.1893 | -100.5072
7.1931 | -100.4704
7.1968 | -100.4447
7.2005 | -100.4292
7.2042 | -100.4161
7.2079 | -100.4118
7.2117 | -100.4165
7.2154 | -100.4282
7.2191 | -100.4463
7.2228 | -100.4685
7.2265 | -100.4956
7.2303 | -100.5233
7.234 | -100.5489
7.2377 | -100.5724
7.2414 | -100.5862
7.2451 | -100.5904
7.2489 | -100.5884
7.2526 | -100.5822
7.2563 | -100.572
7.26 | -100.5477
7.2637 | -100.5217
7.2675 | -100.4968
7.2712 | -100.4764
7.2749 | -100.4664
7.2786 | -100.4597
7.2823 | -100.463
7.286 | -100.4813
7.2898 | -100.5082
7.2935 | -100.5332
7.2972 | -100.5638
7.3009 | -100.592
7.3046 | -100.6127
7.3084 | -100.6373
7.3121 | -100.6619
7.3158 | -100.6816
7.3195 | -100.6997
7.3232 | -100.7136
7.327 | -100.7235
7.3307 | -100.7267
7.3344 | -100.7231
7.3381 | -100.7231
7.3418 | -100.7257
7.3456 | -100.7264
7.3493 | -100.7325
7.353 | -100.7488
7.3567 | -100.7677
7.3604 | -100.7871
7.3642 | -100.8163
7.3679 | -100.8519
7.3716 | -100.8889
7.3753 | -100.9239
7.379 | -100.9574
7.3827 | -100.9878
7.3865 | -101.0121
7.3902 | -101.034
7.3939 | -101.0459
7.3976 | -101.0547
7.4013 | -101.0699
7.4051 | -101.0814
7.4088 | -101.0944
7.4125 | -101.1033
7.4162 | -101.1091
7.4199 | -101.1158
7.4237 | -101.1202
7.4274 | -101.1272
7.4311 | -101.1409
7.4348 | -101.155
7.4385 | -101.1703
7.4423 | -101.1907
7.446 | -101.2063
7.4497 | -101.2149
7.4534 | -101.226
7.4571 | -101.2273
7.4609 | -101.2221
7.4646 | -101.2121
7.4683 | -101.1989
7.472 | -101.172
7.4757 | -101.1315
7.4795 | -101.0846
7.4832 | -101.0295
7.4869 | -100.9635
7.4906 | -100.8897
7.4943 | -100.8173
7.498 | -100.7377
7.5018 | -100.6637
7.5055 | -100.5907
7.5092 | -100.5456
7.5129 | -100.4895
7.5166 | -100.442
7.5204 | -100.4081
7.5241 | -100.3902
7.5278 | -100.382
7.5315 | -100.3837
7.5352 | -100.3956
7.539 | -100.4168
7.5427 | -100.4415
7.5464 | -100.4753
7.5501 | -100.5152
7.5538 | -100.56
7.5576 | -100.609
7.5613 | -100.6634
7.565 | -100.7139
7.5687 | -100.7671
7.5724 | -100.8157
7.5762 | -100.8621
7.5799 | -100.9002
7.5836 | -100.9386
7.5873 | -100.9804
7.591 | -101.0211
7.5947 | -101.0628
7.5985 | -101.1099
7.6022 | -101.1547
7.6059 | -101.1911
7.6096 | -101.2283
7.6133 | -101.2632
7.6171 | -101.3001
7.6208 | -101.3341
7.6245 | -101.3621
7.6282 | -101.38
7.6319 | -101.3974
7.6357 | -101.4101
7.6394 | -101.4124
7.6431 | -101.4063
7.6468 | -101.4009
7.6505 | -101.3873
7.6543 | -101.3608
7.658 | -101.3353
7.6617 | -101.3013
7.6654 | -101.2606
7.6691 | -101.2091
7.6729 | -101.1577
7.6766 | -101.1102
7.6803 | -101.0714
7.684 | -101.0316
7.6877 | -101.0006
7.6914 | -100.9743
7.6952 | -100.9547
7.6989 | -100.943
7.7026 | -100.935
7.7063 | -100.9297
7.71 | -100.9295
7.7138 | -100.9329
7.7175 | -100.9448
7.7212 | -100.9581
7.7249 | -100.9729
7.7286 | -100.989
7.7324 | -101.0057
7.7361 | -101.0143
7.7398 | -101.0241
7.7435 | -101.0361
7.7472 | -101.0412
7.751 | -101.0382
7.7547 | -101.0334
7.7584 | -101.0288
7.7621 | -101.0196
7.7658 | -101.0131
7.7696 | -101.01
7.7733 | -101.0006
7.777 | -100.99
7.7807 | -100.9807
7.7844 | -100.9778
7.7881 | -100.969
7.7919 | -100.9527
7.7956 | -100.9344
7.7993 | -100.9216
7.803 | -100.9086
7.8067 | -100.899
7.8105 | -100.8881
7.8142 | -100.8828
7.8179 | -100.886
7.8216 | -100.9003
7.8253 | -100.9219
7.8291 | -100.9516
7.8328 | -100.9865
7.8365 | -101.0298
7.8402 | -101.0792
7.8439 | -101.133
7.8477 | -101.1896
7.8514 | -101.2534
7.8551 | -101.3181
7.8588 | -101.382
7.8625 | -101.4388
7.8663 | -101.4949
7.87 | -101.5472
7.8737 | -101.5861
7.8774 | -101.6108
7.8811 | -101.6267
7.8849 | -101.6335
7.8886 | -101.6333
7.8923 | -101.6229
7.896 | -101.6033
7.8997 | -101.5698
7.9034 | -101.5393
7.9072 | -101.5055
7.9109 | -101.4664
7.9146 | -101.4292
7.9183 | -101.3946
7.922 | -101.3638
7.9258 | -101.3417
7.9295 | -101.3133
7.9332 | -101.2889
7.9369 | -101.266
7.9406 | -101.24
7.9444 | -101.2034
7.9481 | -101.1774
7.9518 | -101.1487
7.9555 | -101.1145
7.9592 | -101.0803
7.963 | -101.0442
7.9667 | -101.0059
7.9704 | -100.9637
7.9741 | -100.9197
7.9778 | -100.8757
7.9816 | -100.8338
7.9853 | -100.7954
7.989 | -100.7679
7.9927 | -100.7501
7.9964 | -100.7367
8.0001 | -100.7365
8.0039 | -100.7429
8.0076 | -100.7542
8.0113 | -100.7736
8.015 | -100.7986
8.0187 | -100.8211
8.0225 | -100.8446
8.0262 | -100.8622
8.0299 | -100.8664
8.0336 | -100.8641
8.0373 | -100.8487
8.0411 | -100.8177
8.0448 | -100.7866
8.0485 | -100.7454
8.0522 | -100.7157
8.0559 | -100.6902
8.0597 | -100.6729
8.0634 | -100.6637
8.0671 | -100.6478
8.0708 | -100.6427
8.0745 | -100.6466
8.0783 | -100.655
8.082 | -100.6709
8.0857 | -100.6971
8.0894 | -100.7332
8.0931 | -100.7754
8.0968 | -100.8213
8.1006 | -100.8669
8.1043 | -100.9152
8.108 | -100.964
8.1117 | -101.0117
8.1154 | -101.0611
8.1192 | -101.0997
8.1229 | -101.1339
8.1266 | -101.163
8.1303 | -101.1769
8.134 | -101.1821
8.1378 | -101.1798
8.1415 | -101.1707
8.1452 | -101.1575
8.1489 | -101.1449
8.1526 | -101.1338
8.1564 | -101.1271
8.1601 | -101.1261
8.1638 | -101.1213
8.1675 | -101.1216
8.1712 | -101.1274
8.175 | -101.1421
8.1787 | -101.1545
8.1824 | -101.1747
8.1861 | -101.2022
8.1898 | -101.2323
8.1936 | -101.2586
8.1973 | -101.285
8.201 | -101.3061
8.2047 | -101.3259
8.2084 | -101.3462
8.2121 | -101.3655
8.2159 | -101.3852
8.2196 | -101.4014
8.2233 | -101.415
8.227 | -101.4248
8.2307 | -101.4266
8.2345 | -101.4184
8.2382 | -101.4077
8.2419 | -101.3986
8.2456 | -101.3908
8.2493 | -101.3841
8.2531 | -101.374
8.2568 | -101.3578
8.2605 | -101.3429
8.2642 | -101.3274
8.2679 | -101.3097
8.2717 | -101.2862
8.2754 | -101.2605
8.2791 | -101.2424
8.2828 | -101.2257
8.2865 | -101.2048
8.2903 | -101.1842
8.294 | -101.1637
8.2977 | -101.1448
8.3014 | -101.1358
8.3051 | -101.1289
8.3088 | -101.1248
8.3126 | -101.1261
8.3163 | -101.1346
8.32 | -101.146
8.3237 | -101.1647
8.3274 | -101.1884
8.3312 | -101.209
8.3349 | -101.2241
8.3386 | -101.2441
8.3423 | -101.2726
8.346 | -101.3001
8.3498 | -101.3209
8.3535 | -101.3396
8.3572 | -101.3597
8.3609 | -101.3697
8.3646 | -101.3696
8.3684 | -101.3558
8.3721 | -101.3367
8.3758 | -101.3189
8.3795 | -101.2967
8.3832 | -101.2643
8.387 | -101.2312
8.3907 | -101.203
8.3944 | -101.179
8.3981 | -101.1522
8.4018 | -101.1177
8.4055 | -101.0875
8.4093 | -101.0689
8.413 | -101.0537
8.4167 | -101.038
8.4204 | -101.0236
8.4241 | -101.0097
8.4279 | -101.0001
8.4316 | -100.996
8.4353 | -100.998
8.439 | -101.0046
8.4427 | -101.0086
8.4465 | -101.0166
8.4502 | -101.0268
8.4539 | -101.0371
8.4576 | -101.0513
8.4613 | -101.0664
8.4651 | -101.0788
8.4688 | -101.096
8.4725 | -101.1167
8.4762 | -101.1352
8.4799 | -101.1533
8.4837 | -101.172
8.4874 | -101.1939
8.4911 | -101.2217
8.4948 | -101.2478
8.4985 | -101.2735
8.5023 | -101.3053
8.506 | -101.3424
8.5097 | -101.379
8.5134 | -101.418
8.5171 | -101.463
8.5208 | -101.5114
8.5246 | -101.5661
8.5283 | -101.6172
8.532 | -101.6656
8.5357 | -101.7154
8.5394 | -101.767
8.5432 | -101.814
8.5469 | -101.8532
8.5506 | -101.8777
8.5543 | -101.8934
8.558 | -101.8991
8.5618 | -101.8942
8.5655 | -101.882
8.5692 | -101.8562
8.5729 | -101.8189
8.5766 | -101.7733
8.5804 | -101.7225
8.5841 | -101.6585
8.5878 | -101.607
8.5915 | -101.5589
8.5952 | -101.5064
8.599 | -101.4529
8.6027 | -101.4065
8.6064 | -101.3595
8.6101 | -101.3152
8.6138 | -101.2714
8.6175 | -101.2298
8.6213 | -101.1948
8.625 | -101.165
8.6287 | -101.1396
8.6324 | -101.1201
8.6361 | -101.1099
8.6399 | -101.105
8.6436 | -101.1072
8.6473 | -101.1226
8.651 | -101.144
8.6547 | -101.1724
8.6585 | -101.2072
8.6622 | -101.2473
8.6659 | -101.295
8.6696 | -101.3506
8.6733 | -101.4129
8.6771 | -101.4811
8.6808 | -101.5437
8.6845 | -101.6028
8.6882 | -101.6663
8.6919 | -101.7305
8.6957 | -101.7893
8.6994 | -101.8408
8.7031 | -101.8879
8.7068 | -101.9253
8.7105 | -101.9478
8.7142 | -101.9604
8.718 | -101.9677
8.7217 | -101.9684
8.7254 | -101.9693
8.7291 | -101.9683
8.7328 | -101.9557
8.7366 | -101.9416
8.7403 | -101.9269
8.744 | -101.9138
8.7477 | -101.902
8.7514 | -101.8904
8.7552 | -101.8827
8.7589 | -101.8743
8.7626 | -101.8624
8.7663 | -101.8432
8.77 | -101.8258
8.7738 | -101.8069
8.7775 | -101.7853
8.7812 | -101.7676
8.7849 | -101.7454
8.7886 | -101.7178
8.7924 | -101.6866
8.7961 | -101.6522
8.7998 | -101.6143
8.8035 | -101.5768
8.8072 | -101.5388
8.811 | -101.4997
8.8147 | -101.4658
8.8184 | -101.4337
8.8221 | -101.4047
8.8258 | -101.3798
8.8295 | -101.3568
8.8333 | -101.3355
8.837 | -101.3201
8.8407 | -101.3133
8.8444 | -101.3101
8.8481 | -101.3105
8.8519 | -101.3147
8.8556 | -101.3255
8.8593 | -101.3367
8.863 | -101.3542
8.8667 | -101.37
8.8705 | -101.38
8.8742 | -101.3926
8.8779 | -101.4057
8.8816 | -101.4123
8.8853 | -101.4155
8.8891 | -101.4188
8.8928 | -101.4225
8.8965 | -101.4278
8.9002 | -101.4325
8.9039 | -101.4334
8.9077 | -101.433
8.9114 | -101.4338
8.9151 | -101.4371
8.9188 | -101.441
8.9225 | -101.4521
8.9262 | -101.4631
8.93 | -101.4729
8.9337 | -101.4775
8.9374 | -101.4801
8.9411 | -101.4923
8.9448 | -101.4963
8.9486 | -101.4953
8.9523 | -101.4911
8.956 | -101.4843
8.9597 | -101.4749
8.9634 | -101.4613
8.9672 | -101.4425
8.9709 | -101.41
8.9746 | -101.3753
8.9783 | -101.3393
8.982 | -101.3043
8.9858 | -101.2647
8.9895 | -101.2308
8.9932 | -101.1978
8.9969 | -101.1707
9.0006 | -101.1529
9.0044 | -101.1357
9.0081 | -101.1274
9.0118 | -101.1292
9.0155 | -101.1387
9.0192 | -101.1603
9.0229 | -101.1944
9.0267 | -101.2337
9.0304 | -101.2795
9.0341 | -101.3355
9.0378 | -101.3979
9.0415 | -101.4609
9.0453 | -101.5269
9.049 | -101.5929
9.0527 | -101.6511
9.0564 | -101.7078
9.0601 | -101.7629
9.0639 | -101.811
9.0676 | -101.8511
9.0713 | -101.8778
9.075 | -101.8952
9.0787 | -101.9059
9.0825 | -101.9152
9.0862 | -101.9179
9.0899 | -101.9154
9.0936 | -101.9065
9.0973 | -101.8919
9.1011 | -101.8787
9.1048 | -101.8587
9.1085 | -101.8325
9.1122 | -101.8028
9.1159 | -101.7789
9.1196 | -101.7637
9.1234 | -101.7413
9.1271 | -101.7245
9.1308 | -101.7094
9.1345 | -101.6977
9.1382 | -101.6928
9.142 | -101.6856
9.1457 | -101.6857
9.1494 | -101.6895
9.1531 | -101.6915
9.1568 | -101.6963
9.1606 | -101.7055
9.1643 | -101.7106
9.168 | -101.7166
9.1717 | -101.7205
9.1754 | -101.7197
9.1792 | -101.7165
9.1829 | -101.7108
9.1866 | -101.6982
9.1903 | -101.684
9.194 | -101.6683
9.1978 | -101.6542
9.2015 | -101.6404
9.2052 | -101.6237
9.2089 | -101.6092
9.2126 | -101.5938
9.2164 | -101.5772
9.2201 | -101.5675
9.2238 | -101.5575
9.2275 | -101.5474
9.2312 | -101.537
9.2349 | -101.5328
9.2387 | -101.5304
9.2424 | -101.5267
9.2461 | -101.5263
9.2498 | -101.5237
9.2535 | -101.5205
9.2573 | -101.5218
9.261 | -101.5208
9.2647 | -101.5154
9.2684 | -101.5133
9.2721 | -101.5078
9.2759 | -101.5029
9.2796 | -101.5042
9.2833 | -101.5082
9.287 | -101.5118
9.2907 | -101.5146
9.2945 | -101.5198
9.2982 | -101.5296
9.3019 | -101.5438
9.3056 | -101.5589
9.3093 | -101.5766
9.3131 | -101.6001
9.3168 | -101.6239
9.3205 | -101.6488
9.3242 | -101.6752
9.3279 | -101.6964
9.3316 | -101.7148
9.3354 | -101.7333
9.3391 | -101.7477
9.3428 | -101.7629
9.3465 | -101.7769
9.3502 | -101.7954
9.354 | -101.8114
9.3577 | -101.8252
9.3614 | -101.8415
9.3651 | -101.8597
9.3688 | -101.8764
9.3726 | -101.8899
9.3763 | -101.9059
9.38 | -101.9234
9.3837 | -101.9419
9.3874 | -101.9602
9.3912 | -101.9758
9.3949 | -101.9915
9.3986 | -102.0101
9.4023 | -102.0223
9.406 | -102.029
9.4098 | -102.0424
9.4135 | -102.0526
9.4172 | -102.0561
9.4209 | -102.0507
9.4246 | -102.0457
9.4283 | -102.041
9.4321 | -102.0258
9.4358 | -102.0095
9.4395 | -101.991
9.4432 | -101.9689
9.4469 | -101.9456
9.4507 | -101.9202
9.4544 | -101.8956
9.4581 | -101.8723
9.4618 | -101.855
9.4655 | -101.8396
9.4693 | -101.8262
9.473 | -101.8173
9.4767 | -101.8164
9.4804 | -101.8137
9.4841 | -101.8134
9.4879 | -101.8143
9.4916 | -101.8234
9.4953 | -101.8332
9.499 | -101.845
9.5027 | -101.8563
9.5065 | -101.8686
9.5102 | -101.8794
9.5139 | -101.8876
9.5176 | -101.8982
9.5213 | -101.9049
9.525 | -101.9131
9.5288 | -101.9207
9.5325 | -101.9273
9.5362 | -101.9295
9.5399 | -101.936
9.5436 | -101.9425
9.5474 | -101.9411
9.5511 | -101.9387
9.5548 | -101.9418
9.5585 | -101.9425
9.5622 | -101.9439
9.566 | -101.9418
9.5697 | -101.9379
9.5734 | -101.9255
9.5771 | -101.9148
9.5808 | -101.9026
9.5846 | -101.8884
9.5883 | -101.869
9.592 | -101.8458
9.5957 | -101.8237
9.5994 | -101.7991
9.6032 | -101.7788
9.6069 | -101.7599
9.6106 | -101.741
9.6143 | -101.726
9.618 | -101.7138
9.6218 | -101.7025
9.6255 | -101.6906
9.6292 | -101.6814
9.6329 | -101.6784
9.6366 | -101.6791
9.6403 | -101.6831
9.6441 | -101.6816
9.6478 | -101.6802
9.6515 | -101.6787
9.6552 | -101.6714
9.6589 | -101.6687
9.6627 | -101.6679
9.6664 | -101.6638
9.6701 | -101.6563
9.6738 | -101.6521
9.6775 | -101.6535
9.6813 | -101.6543
9.685 | -101.6546
9.6887 | -101.6573
9.6924 | -101.6605
9.6961 | -101.6615
9.6999 | -101.6682
9.7036 | -101.6734
9.7073 | -101.6793
9.711 | -101.6866
9.7147 | -101.6947
9.7185 | -101.7001
9.7222 | -101.71
9.7259 | -101.7191
9.7296 | -101.7286
9.7333 | -101.7396
9.737 | -101.7507
9.7408 | -101.7665
9.7445 | -101.7838
9.7482 | -101.8019
9.7519 | -101.8199
9.7556 | -101.8387
9.7594 | -101.8576
9.7631 | -101.8785
9.7668 | -101.8989
9.7705 | -101.9186
9.7742 | -101.9402
9.778 | -101.9642
9.7817 | -101.9902
9.7854 | -102.0215
9.7891 | -102.0542
9.7928 | -102.0852
9.7966 | -102.121
9.8003 | -102.1563
9.804 | -102.1889
9.8077 | -102.2199
9.8114 | -102.2506
9.8152 | -102.2851
9.8189 | -102.3122
9.8226 | -102.3315
9.8263 | -102.3511
9.83 | -102.3662
9.8337 | -102.376
9.8375 | -102.3854
9.8412 | -102.3904
9.8449 | -102.3914
9.8486 | -102.3913
9.8523 | -102.3905
9.8561 | -102.3882
9.8598 | -102.3854
9.8635 | -102.3815
9.8672 | -102.3745
9.8709 | -102.3727
9.8747 | -102.3705
9.8784 | -102.3658
9.8821 | -102.3624
9.8858 | -102.3578
9.8895 | -102.3512
9.8933 | -102.3454
9.897 | -102.3387
9.9007 | -102.3328
9.9044 | -102.3299
9.9081 | -102.323
9.9119 | -102.3086
9.9156 | -102.2931
9.9193 | -102.2749
9.923 | -102.2496
9.9267 | -102.2096
9.9305 | -102.1521
9.9342 | -102.108
9.9379 | -102.064
9.9416 | -102.0151
9.9453 | -101.9635
9.949 | -101.9056
9.9528 | -101.8472
9.9565 | -101.791
9.9602 | -101.7318
9.9639 | -101.6709
9.9676 | -101.6123
9.9714 | -101.5569
9.9751 | -101.5054
9.9788 | -101.4608
9.9825 | -101.4195
9.9862 | -101.3862
9.99 | -101.3621
9.9937 | -101.3473
9.9974 | -101.3436
10.0011 | -101.3505
10.0048 | -101.3641
10.0086 | -101.3886
10.0123 | -101.4238
10.016 | -101.4628
10.0197 | -101.5052
10.0234 | -101.5535
10.0272 | -101.6043
10.0309 | -101.6538
10.0346 | -101.7024
10.0383 | -101.7456
10.042 | -101.7821
10.0457 | -101.8117
10.0495 | -101.8348
10.0532 | -101.8516
10.0569 | -101.8639
10.0606 | -101.869
10.0643 | -101.8662
10.0681 | -101.864
10.0718 | -101.867
10.0755 | -101.8742
10.0792 | -101.8577
10.0829 | -101.8419
10.0867 | -101.8294
10.0904 | -101.8235
10.0941 | -101.8256
10.0978 | -101.8314
10.1015 | -101.8397
10.1053 | -101.8527
10.109 | -101.8725
10.1127 | -101.8969
10.1164 | -101.925
10.1201 | -101.9585
10.1239 | -101.9947
10.1276 | -102.032
10.1313 | -102.0689
10.135 | -102.1049
10.1387 | -102.1393
10.1424 | -102.1731
10.1462 | -102.2054
10.1499 | -102.2374
10.1536 | -102.2597
10.1573 | -102.2757
10.161 | -102.2931
10.1648 | -102.3113
10.1685 | -102.3218
10.1722 | -102.3256
10.1759 | -102.3256
10.1796 | -102.3207
10.1834 | -102.3165
10.1871 | -102.3106
10.1908 | -102.2983
10.1945 | -102.2775
10.1982 | -102.2604
10.202 | -102.2425
10.2057 | -102.2235
10.2094 | -102.1999
10.2131 | -102.1755
10.2168 | -102.1493
10.2206 | -102.12
10.2243 | -102.0939
10.228 | -102.0659
10.2317 | -102.0383
10.2354 | -102.009
10.2391 | -101.9789
10.2429 | -101.9491
10.2466 | -101.9245
10.2503 | -101.9039
10.254 | -101.8858
10.2577 | -101.8709
10.2615 | -101.8594
10.2652 | -101.8475
10.2689 | -101.8403
10.2726 | -101.8396
10.2763 | -101.8424
10.2801 | -101.8463
10.2838 | -101.8556
10.2875 | -101.8688
10.2912 | -101.8856
10.2949 | -101.9054
10.2987 | -101.9293
10.3024 | -101.9609
10.3061 | -101.993
10.3098 | -102.0242
10.3135 | -102.0596
10.3173 | -102.0983
10.321 | -102.1344
10.3247 | -102.1698
10.3284 | -102.2001
10.3321 | -102.2289
10.3359 | -102.2598
10.3396 | -102.2867
10.3433 | -102.3084
10.347 | -102.3269
10.3507 | -102.3458
10.3544 | -102.3603
10.3582 | -102.3707
10.3619 | -102.3829
10.3656 | -102.3933
10.3693 | -102.4024
10.373 | -102.4126
10.3768 | -102.4152
10.3805 | -102.4172
10.3842 | -102.4202
10.3879 | -102.4229
10.3916 | -102.4239
10.3954 | -102.4246
10.3991 | -102.4227
10.4028 | -102.4246
10.4065 | -102.4293
10.4102 | -102.4335
10.414 | -102.4396
10.4177 | -102.4516
10.4214 | -102.4644
10.4251 | -102.4773
10.4288 | -102.4919
10.4326 | -102.512
10.4363 | -102.5321
10.44 | -102.5521
10.4437 | -102.5703
10.4474 | -102.5884
10.4511 | -102.6015
10.4549 | -102.6107
10.4586 | -102.622
10.4623 | -102.6282
10.466 | -102.6297
10.4697 | -102.6283
10.4735 | -102.6281
10.4772 | -102.631
10.4809 | -102.6389
10.4846 | -102.6404
10.4883 | -102.6374
10.4921 | -102.6369
10.4958 | -102.6397
10.4995 | -102.6472
10.5032 | -102.6522
10.5069 | -102.6594
10.5107 | -102.669
10.5144 | -102.679
10.5181 | -102.6946
10.5218 | -102.7112
10.5255 | -102.7233
10.5293 | -102.7392
10.533 | -102.7591
10.5367 | -102.7748
10.5404 | -102.7869
10.5441 | -102.7994
10.5478 | -102.8068
10.5516 | -102.8088
10.5553 | -102.8067
10.559 | -102.8034
10.5627 | -102.7965
10.5664 | -102.7832
10.5702 | -102.7651
10.5739 | -102.748
10.5776 | -102.73
10.5813 | -102.7137
10.585 | -102.6945
10.5888 | -102.6744
10.5925 | -102.6548
10.5962 | -102.6394
10.5999 | -102.6257
10.6036 | -102.6204
10.6074 | -102.6161
10.6111 | -102.6085
10.6148 | -102.6069
10.6185 | -102.6055
10.6222 | -102.6061
10.626 | -102.6067
10.6297 | -102.6026
10.6334 | -102.5981
10.6371 | -102.5958
10.6408 | -102.5921
10.6445 | -102.5868
10.6483 | -102.5789
10.652 | -102.5687
10.6557 | -102.5598
10.6594 | -102.5528
10.6631 | -102.544
10.6669 | -102.5337
10.6706 | -102.5239
10.6743 | -102.5162
10.678 | -102.5148
10.6817 | -102.5145
10.6855 | -102.5145
10.6892 | -102.5223
10.6929 | -102.5348
10.6966 | -102.5461
10.7003 | -102.5596
10.7041 | -102.576
10.7078 | -102.5925
10.7115 | -102.6093
10.7152 | -102.6257
10.7189 | -102.643
10.7227 | -102.6589
10.7264 | -102.6691
10.7301 | -102.6667
10.7338 | -102.6735
10.7375 | -102.6775
10.7413 | -102.6783
10.745 | -102.6744
10.7487 | -102.6684
10.7524 | -102.6606
10.7561 | -102.6532
10.7598 | -102.6429
10.7636 | -102.6326
10.7673 | -102.6238
10.771 | -102.6142
10.7747 | -102.6091
10.7784 | -102.6037
10.7822 | -102.5991
10.7859 | -102.5997
10.7896 | -102.6038
10.7933 | -102.6109
10.797 | -102.6239
10.8008 | -102.6382
10.8045 | -102.6528
10.8082 | -102.6707
10.8119 | -102.6911
10.8156 | -102.7103
10.8194 | -102.7316
10.8231 | -102.7523
10.8268 | -102.7722
10.8305 | -102.7885
10.8342 | -102.8053
10.838 | -102.8212
10.8417 | -102.8292
10.8454 | -102.8286
10.8491 | -102.8305
10.8528 | -102.8281
10.8565 | -102.8177
10.8603 | -102.8024
10.864 | -102.7837
10.8677 | -102.7635
10.8714 | -102.7418
10.8751 | -102.7191
10.8789 | -102.6915
10.8826 | -102.6641
10.8863 | -102.6401
10.89 | -102.6192
10.8937 | -102.595
10.8975 | -102.5723
10.9012 | -102.5586
10.9049 | -102.5495
10.9086 | -102.5435
10.9123 | -102.5387
10.9161 | -102.5364
10.9198 | -102.5348
10.9235 | -102.5355
10.9272 | -102.5415
10.9309 | -102.5446
10.9347 | -102.5457
10.9384 | -102.5461
10.9421 | -102.5462
10.9458 | -102.5439
10.9495 | -102.5393
10.9532 | -102.531
10.957 | -102.52
10.9607 | -102.5088
10.9644 | -102.4961
10.9681 | -102.4826
10.9718 | -102.4668
10.9756 | -102.4522
10.9793 | -102.4379
10.983 | -102.4237
10.9867 | -102.411
10.9904 | -102.4036
10.9942 | -102.3988
10.9979 | -102.3966
11.0016 | -102.4008
11.0053 | -102.4075
11.009 | -102.4137
11.0128 | -102.4283
11.0165 | -102.4449
11.0202 | -102.4607
11.0239 | -102.4776
11.0276 | -102.4918
11.0314 | -102.5055
11.0351 | -102.5184
11.0388 | -102.5341
11.0425 | -102.5534
11.0462 | -102.5627
11.05 | -102.5697
11.0537 | -102.5837
11.0574 | -102.5974
11.0611 | -102.6063
11.0648 | -102.6178
11.0685 | -102.6297
11.0723 | -102.6421
11.076 | -102.6561
11.0797 | -102.671
11.0834 | -102.6806
11.0871 | -102.6856
11.0909 | -102.6959
11.0946 | -102.7046
11.0983 | -102.7138
11.102 | -102.7215
11.1057 | -102.7272
11.1095 | -102.7276
11.1132 | -102.7289
11.1169 | -102.7289
11.1206 | -102.7246
11.1243 | -102.7206
11.1281 | -102.7168
11.1318 | -102.7137
11.1355 | -102.709
11.1392 | -102.7056
11.1429 | -102.7038
11.1467 | -102.6986
11.1504 | -102.6931
11.1541 | -102.6861
11.1578 | -102.6785
11.1615 | -102.6737
11.1652 | -102.6708
11.169 | -102.6662
11.1727 | -102.6605
11.1764 | -102.6578
11.1801 | -102.6571
11.1838 | -102.6523
11.1876 | -102.6543
11.1913 | -102.6601
11.195 | -102.6641
11.1987 | -102.6661
11.2024 | -102.6669
11.2062 | -102.6725
11.2099 | -102.6775
11.2136 | -102.6793
11.2173 | -102.6813
11.221 | -102.6834
11.2248 | -102.6829
11.2285 | -102.6831
11.2322 | -102.6833
11.2359 | -102.6847
11.2396 | -102.6866
11.2434 | -102.6941
11.2471 | -102.7041
11.2508 | -102.7096
11.2545 | -102.7167
11.2582 | -102.7261
11.2619 | -102.7371
11.2657 | -102.7576
11.2694 | -102.7695
11.2731 | -102.7853
11.2768 | -102.8021
11.2805 | -102.8174
11.2843 | -102.8325
11.288 | -102.8462
11.2917 | -102.8547
11.2954 | -102.8622
11.2991 | -102.8678
11.3029 | -102.8654
11.3066 | -102.8601
11.3103 | -102.8534
11.314 | -102.8417
11.3177 | -102.8291
11.3215 | -102.8151
11.3252 | -102.7957
11.3289 | -102.7722
11.3326 | -102.7469
11.3363 | -102.7187
11.3401 | -102.6879
11.3438 | -102.6598
11.3475 | -102.6366
11.3512 | -102.6152
11.3549 | -102.5937
11.3586 | -102.573
11.3624 | -102.5561
11.3661 | -102.5449
11.3698 | -102.5371
11.3735 | -102.5302
11.3772 | -102.5271
11.381 | -102.5249
11.3847 | -102.5234
11.3884 | -102.5224
11.3921 | -102.5216
11.3958 | -102.5212
11.3996 | -102.5222
11.4033 | -102.5242
11.407 | -102.5239
11.4107 | -102.521
11.4144 | -102.5214
11.4182 | -102.5173
11.4219 | -102.5121
11.4256 | -102.5077
11.4293 | -102.5002
11.433 | -102.4931
11.4368 | -102.4861
11.4405 | -102.4778
11.4442 | -102.4664
11.4479 | -102.4554
11.4516 | -102.4439
11.4553 | -102.4308
11.4591 | -102.4177
11.4628 | -102.4031
11.4665 | -102.3911
11.4702 | -102.3842
11.4739 | -102.3772
11.4777 | -102.3697
11.4814 | -102.3613
11.4851 | -102.3508
11.4888 | -102.3431
11.4925 | -102.3405
11.4963 | -102.3361
11.5 | -102.3336
11.5037 | -102.333
11.5074 | -102.3368
11.5111 | -102.3398
11.5149 | -102.3394
11.5186 | -102.3363
11.5223 | -102.3376
11.526 | -102.343
11.5297 | -102.3494
11.5335 | -102.3509
11.5372 | -102.354
11.5409 | -102.3601
11.5446 | -102.3646
11.5483 | -102.3713
11.5521 | -102.3778
11.5558 | -102.3863
11.5595 | -102.3937
11.5632 | -102.4006
11.5669 | -102.4073
11.5706 | -102.4156
11.5744 | -102.4238
11.5781 | -102.4345
11.5818 | -102.4424
11.5855 | -102.4546
11.5892 | -102.4656
11.593 | -102.4788
11.5967 | -102.4947
11.6004 | -102.5105
11.6041 | -102.5245
11.6078 | -102.5364
11.6116 | -102.551
11.6153 | -102.5667
11.619 | -102.583
11.6227 | -102.593
11.6264 | -102.6064
11.6302 | -102.6212
11.6339 | -102.634
11.6376 | -102.644
11.6413 | -102.6533
11.645 | -102.6649
11.6488 | -102.6812
11.6525 | -102.6982
11.6562 | -102.7125
11.6599 | -102.724
11.6636 | -102.742
11.6673 | -102.7573
11.6711 | -102.7692
11.6748 | -102.7783
11.6785 | -102.7836
11.6822 | -102.7925
11.6859 | -102.7983
11.6897 | -102.7991
11.6934 | -102.7987
11.6971 | -102.7966
11.7008 | -102.7993
11.7045 | -102.8026
11.7083 | -102.8056
11.712 | -102.809
11.7157 | -102.8114
11.7194 | -102.8112
11.7231 | -102.8084
11.7269 | -102.8106
11.7306 | -102.8126
11.7343 | -102.8092
11.738 | -102.8078
11.7417 | -102.8104
11.7455 | -102.8098
11.7492 | -102.8059
11.7529 | -102.8011
11.7566 | -102.7976
11.7603 | -102.7904
11.764 | -102.7816
11.7678 | -102.7733
11.7715 | -102.7661
11.7752 | -102.7611
11.7789 | -102.7563
11.7826 | -102.7524
11.7864 | -102.7449
11.7901 | -102.7428
11.7938 | -102.738
11.7975 | -102.7347
11.8012 | -102.743
11.805 | -102.745
11.8087 | -102.7521
11.8124 | -102.76
11.8161 | -102.7686
11.8198 | -102.7762
11.8236 | -102.787
11.8273 | -102.8004
11.831 | -102.8157
11.8347 | -102.8299
11.8384 | -102.8458
11.8422 | -102.865
11.8459 | -102.8845
11.8496 | -102.9069
11.8533 | -102.9248
11.857 | -102.9437
11.8608 | -102.9638
11.8645 | -102.9801
11.8682 | -102.9943
11.8719 | -103.004
11.8756 | -103.0112
11.8793 | -103.0165
11.8831 | -103.0251
11.8868 | -103.0353
11.8905 | -103.0483
11.8942 | -103.0589
11.8979 | -103.066
11.9017 | -103.0746
11.9054 | -103.086
11.9091 | -103.0955
11.9128 | -103.0977
11.9165 | -103.1006
11.9203 | -103.1071
11.924 | -103.111
11.9277 | -103.112
11.9314 | -103.1122
11.9351 | -103.1132
11.9389 | -103.1085
11.9426 | -103.1052
11.9463 | -103.1022
11.95 | -103.0979
11.9537 | -103.0937
11.9575 | -103.0882
11.9612 | -103.079
11.9649 | -103.071
11.9686 | -103.0606
11.9723 | -103.0497
11.976 | -103.0386
11.9798 | -103.0285
11.9835 | -103.0173
11.9872 | -103.0072
11.9909 | -102.9973
11.9946 | -102.9883
11.9984 | -102.9792
12.0021 | -102.9696
12.0058 | -102.9622
12.0095 | -102.9581
12.0132 | -102.9498
12.017 | -102.9433
12.0207 | -102.9407
12.0244 | -102.9383
12.0281 | -102.94
12.0318 | -102.9392
12.0356 | -102.9364
12.0393 | -102.9341
12.043 | -102.9278
12.0467 | -102.9185
12.0504 | -102.9088
12.0542 | -102.901
12.0579 | -102.8866
12.0616 | -102.8697
12.0653 | -102.8558
12.069 | -102.8423
12.0727 | -102.8247
12.0765 | -102.8084
12.0802 | -102.7918
12.0839 | -102.7732
12.0876 | -102.7612
12.0913 | -102.7496
12.0951 | -102.7375
12.0988 | -102.7266
12.1025 | -102.7178
12.1062 | -102.7102
12.1099 | -102.7036
12.1137 | -102.7021
12.1174 | -102.7032
12.1211 | -102.7066
12.1248 | -102.7118
12.1285 | -102.7188
12.1323 | -102.7354
12.136 | -102.756
12.1397 | -102.7775
12.1434 | -102.8044
12.1471 | -102.8321
12.1509 | -102.8651
12.1546 | -102.8985
12.1583 | -102.9326
12.162 | -102.9688
12.1657 | -103.0085
12.1694 | -103.0507
12.1732 | -103.0929
12.1769 | -103.1317
12.1806 | -103.1682
12.1843 | -103.2009
12.188 | -103.2342
12.1918 | -103.2578
12.1955 | -103.275
12.1992 | -103.2888
12.2029 | -103.2999
12.2066 | -103.3107
12.2104 | -103.3173
12.2141 | -103.3199
12.2178 | -103.3224
12.2215 | -103.3246
12.2252 | -103.3215
12.229 | -103.311
12.2327 | -103.3014
12.2364 | -103.2932
12.2401 | -103.2806
12.2438 | -103.2663
12.2476 | -103.2545
12.2513 | -103.2406
12.255 | -103.2304
12.2587 | -103.2248
12.2624 | -103.2138
12.2662 | -103.2034
12.2699 | -103.1966
12.2736 | -103.1922
12.2773 | -103.1871
12.281 | -103.1833
12.2847 | -103.1809
12.2885 | -103.1785
12.2922 | -103.1756
12.2959 | -103.1719
12.2996 | -103.1688
12.3033 | -103.1682
12.3071 | -103.1645
12.3108 | -103.1594
12.3145 | -103.1542
12.3182 | -103.1516
12.3219 | -103.1504
12.3257 | -103.1482
12.3294 | -103.1454
12.3331 | -103.1421
12.3368 | -103.1394
12.3405 | -103.1376
12.3443 | -103.1349
12.348 | -103.1315
12.3517 | -103.1299
12.3554 | -103.1312
12.3591 | -103.1309
12.3629 | -103.131
12.3666 | -103.1348
12.3703 | -103.1411
12.374 | -103.1455
12.3777 | -103.1484
12.3814 | -103.1465
12.3852 | -103.1439
12.3889 | -103.1439
12.3926 | -103.1442
12.3963 | -103.1427
12.4 | -103.1361
12.4038 | -103.1274
12.4075 | -103.1222
12.4112 | -103.1194
12.4149 | -103.1113
12.4186 | -103.1022
12.4224 | -103.0959
12.4261 | -103.0886
12.4298 | -103.0796
12.4335 | -103.0677
12.4372 | -103.0558
12.441 | -103.0495
12.4447 | -103.0434
12.4484 | -103.032
12.4521 | -103.0218
12.4558 | -103.0127
12.4596 | -103.0041
12.4633 | -102.9976
12.467 | -102.9943
12.4707 | -102.9877
12.4744 | -102.9838
12.4781 | -102.9857
12.4819 | -102.9896
12.4856 | -102.9935
12.4893 | -102.9997
12.493 | -103.0076
12.4967 | -103.0142
12.5005 | -103.0239
12.5042 | -103.037
12.5079 | -103.0546
12.5116 | -103.0734
12.5153 | -103.0923
12.5191 | -103.1114
12.5228 | -103.1317
12.5265 | -103.1533
12.5302 | -103.176
12.5339 | -103.1975
12.5377 | -103.2184
12.5414 | -103.2424
12.5451 | -103.2674
12.5488 | -103.2916
12.5525 | -103.3155
12.5563 | -103.3395
12.56 | -103.3666
12.5637 | -103.3946
12.5674 | -103.4191
12.5711 | -103.44
12.5748 | -103.4602
12.5786 | -103.4813
12.5823 | -103.5015
12.586 | -103.5237
12.5897 | -103.5397
12.5934 | -103.5484
12.5972 | -103.5552
12.6009 | -103.5589
12.6046 | -103.5587
12.6083 | -103.559
12.612 | -103.5565
12.6158 | -103.5512
12.6195 | -103.5401
12.6232 | -103.5262
12.6269 | -103.5135
12.6306 | -103.5039
12.6344 | -103.4923
12.6381 | -103.4763
12.6418 | -103.4563
12.6455 | -103.4339
12.6492 | -103.4123
12.653 | -103.3953
12.6567 | -103.376
12.6604 | -103.3548
12.6641 | -103.3314
12.6678 | -103.3102
12.6716 | -103.2917
12.6753 | -103.2752
12.679 | -103.2571
12.6827 | -103.239
12.6864 | -103.2205
12.6901 | -103.2027
12.6939 | -103.1897
12.6976 | -103.1809
12.7013 | -103.176
12.705 | -103.1729
12.7087 | -103.1732
12.7125 | -103.1731
12.7162 | -103.1763
12.7199 | -103.1843
12.7236 | -103.1939
12.7273 | -103.2083
12.7311 | -103.225
12.7348 | -103.2468
12.7385 | -103.2716
12.7422 | -103.2955
12.7459 | -103.3219
12.7497 | -103.3488
12.7534 | -103.376
12.7571 | -103.4027
12.7608 | -103.4298
12.7645 | -103.456
12.7683 | -103.4775
12.772 | -103.4965
12.7757 | -103.5131
12.7794 | -103.5255
12.7831 | -103.5356
12.7868 | -103.5415
12.7906 | -103.5408
12.7943 | -103.537
12.798 | -103.5329
12.8017 | -103.5274
12.8054 | -103.5137
12.8092 | -103.497
12.8129 | -103.474
12.8166 | -103.4487
12.8203 | -103.4264
12.824 | -103.4052
12.8278 | -103.3802
12.8315 | -103.3485
12.8352 | -103.313
12.8389 | -103.2827
12.8426 | -103.2537
12.8464 | -103.2224
12.8501 | -103.1932
12.8538 | -103.1635
12.8575 | -103.1347
12.8612 | -103.1127
12.865 | -103.0914
12.8687 | -103.0743
12.8724 | -103.0616
12.8761 | -103.0601
12.8798 | -103.0494
12.8835 | -103.0399
12.8873 | -103.035
12.891 | -103.0322
12.8947 | -103.0285
12.8984 | -103.0326
12.9021 | -103.0388
12.9059 | -103.0379
12.9096 | -103.0409
12.9133 | -103.0437
12.917 | -103.0498
12.9207 | -103.056
12.9245 | -103.0629
12.9282 | -103.0688
12.9319 | -103.0773
12.9356 | -103.0874
12.9393 | -103.0948
12.9431 | -103.1021
12.9468 | -103.1065
12.9505 | -103.1113
12.9542 | -103.1138
12.9579 | -103.116
12.9617 | -103.1173
12.9654 | -103.1189
12.9691 | -103.119
12.9728 | -103.1202
12.9765 | -103.1212
12.9803 | -103.1197
12.984 | -103.115
12.9877 | -103.1133
12.9914 | -103.1159
12.9951 | -103.1153
12.9988 | -103.1169
13.0026 | -103.1209
13.0063 | -103.1262
13.01 | -103.1277
13.0137 | -103.1322
13.0174 | -103.1428
13.0212 | -103.1567
13.0249 | -103.1662
13.0286 | -103.1751
13.0323 | -103.1879
13.036 | -103.2006
13.0398 | -103.2124
13.0435 | -103.2247
13.0472 | -103.2375
13.0509 | -103.2496
13.0546 | -103.2631
13.0584 | -103.2766
13.0621 | -103.288
13.0658 | -103.3045
13.0695 | -103.3216
13.0732 | -103.3396
13.077 | -103.3565
13.0807 | -103.3737
13.0844 | -103.3942
13.0881 | -103.411
13.0918 | -103.4276
13.0955 | -103.4535
13.0993 | -103.475
13.103 | -103.4915
13.1067 | -103.5048
13.1104 | -103.5183
13.1141 | -103.5313
13.1179 | -103.5473
13.1216 | -103.5598
13.1253 | -103.5703
13.129 | -103.5796
13.1327 | -103.5852
13.1365 | -103.5903
13.1402 | -103.5934
13.1439 | -103.5953
13.1476 | -103.5937
13.1513 | -103.5918
13.1551 | -103.5873
13.1588 | -103.5786
13.1625 | -103.5662
13.1662 | -103.5535
13.1699 | -103.5423
13.1737 | -103.5313
13.1774 | -103.5153
13.1811 | -103.4962
13.1848 | -103.477
13.1885 | -103.4615
13.1922 | -103.4428
13.196 | -103.4191
13.1997 | -103.3981
13.2034 | -103.3774
13.2071 | -103.352
13.2108 | -103.324
13.2146 | -103.3003
13.2183 | -103.2746
13.222 | -103.2514
13.2257 | -103.2306
13.2294 | -103.2095
13.2332 | -103.1873
13.2369 | -103.1614
13.2406 | -103.1378
13.2443 | -103.1139
13.248 | -103.0916
13.2518 | -103.0732
13.2555 | -103.0512
13.2592 | -103.0288
13.2629 | -103.0062
13.2666 | -102.9836
13.2704 | -102.9605
13.2741 | -102.9401
13.2778 | -102.923
13.2815 | -102.9067
13.2852 | -102.8899
13.2889 | -102.8744
13.2927 | -102.8652
13.2964 | -102.8587
13.3001 | -102.8558
13.3038 | -102.8561
13.3075 | -102.8582
13.3113 | -102.8632
13.315 | -102.8692
13.3187 | -102.879
13.3224 | -102.891
13.3261 | -102.9074
13.3299 | -102.9252
13.3336 | -102.9452
13.3373 | -102.9676
13.341 | -102.9909
13.3447 | -103.0163
13.3485 | -103.0426
13.3522 | -103.0646
13.3559 | -103.087
13.3596 | -103.1116
13.3633 | -103.1366
13.3671 | -103.1611
13.3708 | -103.1868
13.3745 | -103.2141
13.3782 | -103.2381
13.3819 | -103.2594
13.3857 | -103.2788
13.3894 | -103.299
13.3931 | -103.314
13.3968 | -103.3282
13.4005 | -103.3409
13.4042 | -103.3485
13.408 | -103.3566
13.4117 | -103.3544
13.4154 | -103.3573
13.4191 | -103.3558
13.4228 | -103.3547
13.4266 | -103.3538
13.4303 | -103.3536
13.434 | -103.3548
13.4377 | -103.3531
13.4414 | -103.3471
13.4452 | -103.3431
13.4489 | -103.3406
13.4526 | -103.3375
13.4563 | -103.3359
13.46 | -103.3376
13.4638 | -103.3432
13.4675 | -103.3443
13.4712 | -103.3417
13.4749 | -103.3382
13.4786 | -103.338
13.4824 | -103.3384
13.4861 | -103.3359
13.4898 | -103.3342
13.4935 | -103.3321
13.4972 | -103.332
13.5009 | -103.3316
13.5047 | -103.3317
13.5084 | -103.3328
13.5121 | -103.3335
13.5158 | -103.3349
13.5195 | -103.3324
13.5233 | -103.3314
13.527 | -103.3315
13.5307 | -103.3296
13.5344 | -103.3293
13.5381 | -103.329
13.5419 | -103.3322
13.5456 | -103.3428
13.5493 | -103.3512
13.553 | -103.359
13.5567 | -103.3681
13.5605 | -103.3806
13.5642 | -103.3916
13.5679 | -103.3991
13.5716 | -103.4074
13.5753 | -103.4203
13.5791 | -103.44
13.5828 | -103.4579
13.5865 | -103.4748
13.5902 | -103.494
13.5939 | -103.5123
13.5976 | -103.5275
13.6014 | -103.5401
13.6051 | -103.5494
13.6088 | -103.5582
13.6125 | -103.5664
13.6162 | -103.573
13.62 | -103.5766
13.6237 | -103.5763
13.6274 | -103.575
13.6311 | -103.5665
13.6348 | -103.5572
13.6386 | -103.5504
13.6423 | -103.5424
13.646 | -103.5334
13.6497 | -103.5237
13.6534 | -103.5134
13.6572 | -103.5001
13.6609 | -103.485
13.6646 | -103.4725
13.6683 | -103.464
13.672 | -103.4575
13.6758 | -103.4508
13.6795 | -103.4447
13.6832 | -103.4418
13.6869 | -103.4434
13.6906 | -103.4443
13.6944 | -103.4424
13.6981 | -103.444
13.7018 | -103.4514
13.7055 | -103.4631
13.7092 | -103.4758
13.7129 | -103.4882
13.7167 | -103.5037
13.7204 | -103.5183
13.7241 | -103.5339
13.7278 | -103.5498
13.7315 | -103.5645
13.7353 | -103.5808
13.739 | -103.5961
13.7427 | -103.6101
13.7464 | -103.6258
13.7501 | -103.6398
13.7539 | -103.6527
13.7576 | -103.6637
13.7613 | -103.6752
13.765 | -103.6871
13.7687 | -103.6974
13.7725 | -103.7056
13.7762 | -103.7105
13.7799 | -103.7149
13.7836 | -103.7171
13.7873 | -103.7143
13.7911 | -103.7107
13.7948 | -103.7093
13.7985 | -103.7077
13.8022 | -103.7039
13.8059 | -103.7027
13.8096 | -103.7008
13.8134 | -103.6964
13.8171 | -103.6925
13.8208 | -103.6903
13.8245 | -103.6867
13.8282 | -103.686
13.832 | -103.6828
13.8357 | -103.679
13.8394 | -103.6716
13.8431 | -103.6671
13.8468 | -103.6628
13.8506 | -103.6589
13.8543 | -103.6576
13.858 | -103.6567
13.8617 | -103.6549
13.8654 | -103.6488
13.8692 | -103.6447
13.8729 | -103.6388
13.8766 | -103.6319
13.8803 | -103.627
13.884 | -103.6186
13.8878 | -103.6133
13.8915 | -103.6144
13.8952 | -103.6169
13.8989 | -103.6174
13.9026 | -103.6094
13.9063 | -103.6047
13.9101 | -103.605
13.9138 | -103.6068
13.9175 | -103.6116
13.9212 | -103.6157
13.9249 | -103.6215
13.9287 | -103.6269
13.9324 | -103.6278
13.9361 | -103.6295
13.9398 | -103.6336
13.9435 | -103.6383
13.9473 | -103.644
13.951 | -103.6489
13.9547 | -103.6514
13.9584 | -103.6527
13.9621 | -103.6553
13.9659 | -103.6583
13.9696 | -103.6603
13.9733 | -103.6611
13.977 | -103.6604
13.9807 | -103.6612
13.9845 | -103.6662
13.9882 | -103.6732
13.9919 | -103.6791
13.9956 | -103.6841
13.9993 | -103.688
14.003 | -103.6933
14.0068 | -103.6929
14.0105 | -103.6921
14.0142 | -103.6955
14.0179 | -103.6974
14.0216 | -103.698
14.0254 | -103.7002
14.0291 | -103.7016
14.0328 | -103.7053
14.0365 | -103.7046
14.0402 | -103.7024
14.044 | -103.6963
14.0477 | -103.6909
14.0514 | -103.685
14.0551 | -103.6803
14.0588 | -103.674
14.0626 | -103.6682
14.0663 | -103.6632
14.07 | -103.6529
14.0737 | -103.6401
14.0774 | -103.632
14.0812 | -103.6247
14.0849 | -103.6151
14.0886 | -103.6007
14.0923 | -103.585
14.096 | -103.5716
14.0998 | -103.5593
14.1035 | -103.5529
14.1072 | -103.5435
14.1109 | -103.5333
14.1146 | -103.5229
14.1183 | -103.5132
14.1221 | -103.5057
14.1258 | -103.4963
14.1295 | -103.4897
14.1332 | -103.4872
14.1369 | -103.4853
14.1407 | -103.4833
14.1444 | -103.4837
14.1481 | -103.4849
14.1518 | -103.4869
14.1555 | -103.4905
14.1593 | -103.499
14.163 | -103.5094
14.1667 | -103.5188
14.1704 | -103.5293
14.1741 | -103.5395
14.1779 | -103.5512
14.1816 | -103.5643
14.1853 | -103.579
14.189 | -103.5951
14.1927 | -103.6121
14.1965 | -103.6335
14.2002 | -103.6534
14.2039 | -103.675
14.2076 | -103.6949
14.2113 | -103.7165
14.215 | -103.735
14.2188 | -103.7518
14.2225 | -103.77
14.2262 | -103.7884
14.2299 | -103.8038
14.2336 | -103.8184
14.2374 | -103.8278
14.2411 | -103.8393
14.2448 | -103.8499
14.2485 | -103.8578
14.2522 | -103.863
14.256 | -103.8652
14.2597 | -103.8633
14.2634 | -103.8601
14.2671 | -103.8528
14.2708 | -103.8474
14.2746 | -103.8422
14.2783 | -103.8371
14.282 | -103.8275
14.2857 | -103.815
14.2894 | -103.8029
14.2932 | -103.7957
14.2969 | -103.7842
14.3006 | -103.7699
14.3043 | -103.7563
14.308 | -103.7447
14.3117 | -103.7326
14.3155 | -103.7156
14.3192 | -103.7002
14.3229 | -103.6855
14.3266 | -103.6685
14.3303 | -103.6561
14.3341 | -103.6439
14.3378 | -103.6309
14.3415 | -103.6148
14.3452 | -103.5984
14.3489 | -103.5844
14.3527 | -103.5726
14.3564 | -103.5637
14.3601 | -103.5525
14.3638 | -103.5412
14.3675 | -103.5306
14.3713 | -103.5223
14.375 | -103.5149
14.3787 | -103.511
14.3824 | -103.5063
14.3861 | -103.5026
14.3899 | -103.5
14.3936 | -103.4991
14.3973 | -103.4992
14.401 | -103.5028
14.4047 | -103.5101
14.4085 | -103.5182
14.4122 | -103.529
14.4159 | -103.5399
14.4196 | -103.5536
14.4233 | -103.5686
14.427 | -103.5848
14.4308 | -103.6019
14.4345 | -103.6209
14.4382 | -103.6398
14.4419 | -103.6607
14.4456 | -103.6792
14.4494 | -103.6993
14.4531 | -103.72
14.4568 | -103.7384
14.4605 | -103.758
14.4642 | -103.7776
14.468 | -103.7966
14.4717 | -103.8158
14.4754 | -103.8296
14.4791 | -103.8412
14.4828 | -103.8503
14.4866 | -103.8463
14.4903 | -103.8501
14.494 | -103.8528
14.4977 | -103.8523
14.5014 | -103.8486
14.5052 | -103.8448
14.5089 | -103.8376
14.5126 | -103.8288
14.5163 | -103.8173
14.52 | -103.8036
14.5237 | -103.7903
14.5275 | -103.7742
14.5312 | -103.7565
14.5349 | -103.7405
14.5386 | -103.7241
14.5423 | -103.7055
14.5461 | -103.6847
14.5498 | -103.668
14.5535 | -103.6542
14.5572 | -103.6378
14.5609 | -103.625
14.5647 | -103.6138
14.5684 | -103.6061
14.5721 | -103.5977
14.5758 | -103.5922
14.5795 | -103.5889
14.5833 | -103.5893
14.587 | -103.5916
14.5907 | -103.595
14.5944 | -103.6029
14.5981 | -103.6122
14.6019 | -103.6232
14.6056 | -103.6343
14.6093 | -103.6433
14.613 | -103.6558
14.6167 | -103.6728
14.6204 | -103.6896
14.6242 | -103.709
14.6279 | -103.7274
14.6316 | -103.7444
14.6353 | -103.761
14.639 | -103.7772
14.6428 | -103.7913
14.6465 | -103.8055
14.6502 | -103.8226
14.6539 | -103.8376
14.6576 | -103.8486
14.6614 | -103.86
14.6651 | -103.8715
14.6688 | -103.8795
14.6725 | -103.8855
14.6762 | -103.8912
14.68 | -103.898
14.6837 | -103.901
14.6874 | -103.9021
14.6911 | -103.902
14.6948 | -103.9009
14.6986 | -103.9006
14.7023 | -103.9013
14.706 | -103.8965
14.7097 | -103.8933
14.7134 | -103.8899
14.7172 | -103.8837
14.7209 | -103.8761
14.7246 | -103.8684
14.7283 | -103.8619
14.732 | -103.8536
14.7357 | -103.8492
14.7395 | -103.8455
14.7432 | -103.8407
14.7469 | -103.8347
14.7506 | -103.8235
14.7543 | -103.8148
14.7581 | -103.8125
14.7618 | -103.8087
14.7655 | -103.8049
14.7692 | -103.8039
14.7729 | -103.8017
14.7767 | -103.7993
14.7804 | -103.7998
14.7841 | -103.8011
14.7878 | -103.7999
14.7915 | -103.7993
14.7953 | -103.7978
14.799 | -103.7983
14.8027 | -103.8036
14.8064 | -103.8078
14.8101 | -103.8116
14.8139 | -103.816
14.8176 | -103.8235
14.8213 | -103.8287
14.825 | -103.8348
14.8287 | -103.8393
14.8324 | -103.8456
14.8362 | -103.8513
14.8399 | -103.8591
14.8436 | -103.8692
14.8473 | -103.8792
14.851 | -103.8877
14.8548 | -103.8989
14.8585 | -103.9091
14.8622 | -103.9164
14.8659 | -103.9281
14.8696 | -103.939
14.8734 | -103.9466
14.8771 | -103.9538
14.8808 | -103.9624
14.8845 | -103.9673
14.8882 | -103.9701
14.892 | -103.9719
14.8957 | -103.9767
14.8994 | -103.9815
14.9031 | -103.9862
14.9068 | -103.9891
14.9106 | -103.989
14.9143 | -103.9863
14.918 | -103.9864
14.9217 | -103.9864
14.9254 | -103.9839
14.9291 | -103.9813
14.9329 | -103.9785
14.9366 | -103.9737
14.9403 | -103.9688
14.944 | -103.9627
14.9477 | -103.9519
14.9515 | -103.9411
14.9552 | -103.9321
14.9589 | -103.9215
14.9626 | -103.9183
14.9663 | -103.9148
14.9701 | -103.9091
14.9738 | -103.9015
14.9775 | -103.8943
14.9812 | -103.887
14.9849 | -103.8809
14.9887 | -103.8778
14.9924 | -103.8754
14.9961 | -103.8754
14.9998 | -103.8792
15.0035 | -103.882
15.0073 | -103.8835
15.011 | -103.8823
15.0147 | -103.8826
15.0184 | -103.886
15.0221 | -103.8963
15.0258 | -103.8962
15.0296 | -103.8968
15.0333 | -103.9003
15.037 | -103.9075
15.0407 | -103.9123
15.0444 | -103.9162
15.0482 | -103.9226
15.0519 | -103.9282
15.0556 | -103.9299
15.0593 | -103.9282
15.063 | -103.9248
15.0668 | -103.9236
15.0705 | -103.9189
15.0742 | -103.9158
15.0779 | -103.9132
15.0816 | -103.9054
15.0854 | -103.8978
15.0891 | -103.8896
15.0928 | -103.8796
15.0965 | -103.8664
15.1002 | -103.8562
15.104 | -103.8465
15.1077 | -103.8343
15.1114 | -103.8184
15.1151 | -103.8018
15.1188 | -103.7858
15.1226 | -103.7719
15.1263 | -103.7594
15.13 | -103.7475
15.1337 | -103.7361
15.1374 | -103.7259
15.1411 | -103.7165
15.1449 | -103.7085
15.1486 | -103.7009
15.1523 | -103.6987
15.156 | -103.6975
15.1597 | -103.6981
15.1635 | -103.7
15.1672 | -103.7035
15.1709 | -103.7091
15.1746 | -103.7185
15.1783 | -103.7247
15.1821 | -103.7297
15.1858 | -103.7362
15.1895 | -103.7455
15.1932 | -103.7549
15.1969 | -103.7652
15.2007 | -103.7735
15.2044 | -103.7774
15.2081 | -103.7825
15.2118 | -103.7905
15.2155 | -103.7957
15.2193 | -103.7993
15.223 | -103.8037
15.2267 | -103.8064
15.2304 | -103.8066
15.2341 | -103.8039
15.2378 | -103.8004
15.2416 | -103.7988
15.2453 | -103.7983
15.249 | -103.7939
15.2527 | -103.787
15.2564 | -103.7784
15.2602 | -103.7698
15.2639 | -103.7625
15.2676 | -103.7526
15.2713 | -103.745
15.275 | -103.7377
15.2788 | -103.7295
15.2825 | -103.7229
15.2862 | -103.718
15.2899 | -103.7163
15.2936 | -103.7124
15.2974 | -103.7093
15.3011 | -103.7106
15.3048 | -103.7128
15.3085 | -103.7174
15.3122 | -103.7231
15.316 | -103.7314
15.3197 | -103.7401
15.3234 | -103.7492
15.3271 | -103.758
15.3308 | -103.7708
15.3345 | -103.7849
15.3383 | -103.7982
15.342 | -103.8081
15.3457 | -103.8179
15.3494 | -103.8305
15.3531 | -103.8414
15.3569 | -103.8497
15.3606 | -103.8573
15.3643 | -103.8629
15.368 | -103.8694
15.3717 | -103.8743
15.3755 | -103.8796
15.3792 | -103.8862
15.3829 | -103.8912
15.3866 | -103.8918
15.3903 | -103.8884
15.3941 | -103.8865
15.3978 | -103.8843
15.4015 | -103.8834
15.4052 | -103.8817
15.4089 | -103.8764
15.4127 | -103.8695
15.4164 | -103.8636
15.4201 | -103.8589
15.4238 | -103.8574
15.4275 | -103.8524
15.4312 | -103.8422
15.435 | -103.8358
15.4387 | -103.8305
15.4424 | -103.822
15.4461 | -103.8163
15.4498 | -103.8098
15.4536 | -103.8029
15.4573 | -103.7961
15.461 | -103.7891
15.4647 | -103.7816
15.4684 | -103.7748
15.4722 | -103.7662
15.4759 | -103.757
15.4796 | -103.7457
15.4833 | -103.7349
15.487 | -103.7248
15.4908 | -103.7141
15.4945 | -103.7058
15.4982 | -103.6981
15.5019 | -103.6886
15.5056 | -103.6777
15.5094 | -103.6669
15.5131 | -103.4642
15.5168 | -102.4826
15.5205 | -102.0186
15.5242 | -101.7519
15.528 | -101.464
15.5317 | -101.1584
15.5354 | -100.8377
15.5391 | -100.5033
15.5428 | -100.1591
15.5465 | -99.8094
15.5503 | -99.456
15.554 | -99.1029
15.5577 | -98.6481
15.5614 | -98.3231
15.5651 | -98.0075
15.5689 | -97.7031
15.5726 | -97.4116
15.5763 | -97.1348
15.58 | -96.8736
15.5837 | -96.6296
15.5875 | -96.4043
15.5912 | -96.1977
15.5949 | -96.0103
15.5986 | -95.8432
15.6023 | -95.6975
15.6061 | -95.5738
15.6098 | -95.4721
15.6135 | -95.393
15.6172 | -95.3364
15.6209 | -95.3027
15.6247 | -95.292
15.6284 | -95.304
15.6321 | -95.3392
15.6358 | -95.3974
15.6395 | -95.478
15.6432 | -95.5805
15.647 | -95.7052
15.6507 | -95.8523
15.6544 | -96.0206
15.6581 | -96.2093
15.6618 | -96.4182
15.6656 | -96.6472
15.6693 | -96.8947
15.673 | -97.16
15.6767 | -97.4416
15.6804 | -97.7382
15.6842 | -98.0481
15.6879 | -98.3698
15.6916 | -98.7016
15.6953 | -99.0408
15.699 | -99.3852
15.7028 | -99.7314
15.7065 | -100.0758
15.7102 | -100.4164
15.7139 | -100.7484
15.7176 | -101.0683
15.7214 | -101.3746
15.7251 | -101.6655
15.7288 | -101.9377
15.7325 | -102.1886
15.7362 | -102.5377
15.7399 | -103.5023
15.7437 | -103.8368
15.7474 | -103.8246
15.7511 | -103.8136
15.7548 | -103.8047
15.7585 | -103.793
15.7623 | -103.7845
15.766 | -103.7783
15.7697 | -103.773
15.7734 | -103.77
15.7771 | -103.7661
15.7809 | -103.7636
15.7846 | -103.765
15.7883 | -103.7687
15.792 | -103.7734
15.7957 | -103.7786
15.7995 | -103.784
15.8032 | -103.7927
15.8069 | -103.8025
15.8106 | -103.8111
15.8143 | -103.8188
15.8181 | -103.8245
15.8218 | -103.8306
15.8255 | -103.8376
15.8292 | -103.8426
15.8329 | -103.8468
15.8367 | -103.8546
15.8404 | -103.8605
15.8441 | -103.8634
15.8478 | -103.863
15.8515 | -103.8589
15.8552 | -103.8559
15.859 | -103.8495
15.8627 | -103.8386
15.8664 | -103.8261
15.8701 | -103.8155
15.8738 | -103.7986
15.8776 | -103.7773
15.8813 | -103.7574
15.885 | -103.7356
15.8887 | -103.7098
15.8924 | -103.6826
15.8962 | -103.6546
15.8999 | -103.6253
15.9036 | -103.5955
15.9073 | -103.5657
15.911 | -103.5382
15.9148 | -103.5132
15.9185 | -103.4887
15.9222 | -103.4662
15.9259 | -103.4453
15.9296 | -103.4268
15.9334 | -103.4107
15.9371 | -103.3966
15.9408 | -103.3866
15.9445 | -103.3815
15.9482 | -103.378
15.9519 | -103.3767
15.9557 | -103.3769
15.9594 | -103.3807
15.9631 | -103.3881
15.9668 | -103.3988
15.9705 | -103.4115
15.9743 | -103.4298
15.978 | -103.4511
15.9817 | -103.4734
15.9854 | -103.4988
15.9891 | -103.5243
15.9929 | -103.5521
15.9966 | -103.5828
16.0003 | -103.6108
16.004 | -103.6371
16.0077 | -103.6659
16.0115 | -103.6943
16.0152 | -103.722
16.0189 | -103.7515
16.0226 | -103.7777
16.0263 | -103.8022
16.0301 | -103.8213
16.0338 | -103.8411
16.0375 | -103.8585
16.0412 | -103.8756
16.0449 | -103.8914
16.0486 | -103.9003
16.0524 | -103.9087
16.0561 | -103.9142
16.0598 | -103.9163
16.0635 | -103.9146
16.0672 | -103.908
16.071 | -103.9008
16.0747 | -103.8916
16.0784 | -103.8766
16.0821 | -103.8606
16.0858 | -103.8465
16.0896 | -103.8327
16.0933 | -103.8187
16.097 | -103.7921
16.1007 | -103.7723
16.1044 | -103.7524
16.1082 | -103.7317
16.1119 | -103.7097
16.1156 | -103.6877
16.1193 | -103.6666
16.123 | -103.6448
16.1268 | -103.6216
16.1305 | -103.5964
16.1342 | -103.5722
16.1379 | -103.549
16.1416 | -103.5241
16.1453 | -103.5008
16.1491 | -103.479
16.1528 | -103.4588
16.1565 | -103.4398
16.1602 | -103.4244
16.1639 | -103.4109
16.1677 | -103.3967
16.1714 | -103.3834
16.1751 | -103.3717
16.1788 | -103.3627
16.1825 | -103.3534
16.1863 | -103.3462
16.19 | -103.3422
16.1937 | -103.3407
16.1974 | -103.3393
16.2011 | -103.3396
16.2049 | -103.3431
16.2086 | -103.3507
16.2123 | -103.3608
16.216 | -103.372
16.2197 | -103.3837
16.2235 | -103.3998
16.2272 | -103.4166
16.2309 | -103.437
16.2346 | -103.46
16.2383 | -103.481
16.242 | -103.5041
16.2458 | -103.5268
16.2495 | -103.5471
16.2532 | -103.5693
16.2569 | -103.591
16.2606 | -103.6151
16.2644 | -103.6379
16.2681 | -103.6597
16.2718 | -103.6778
16.2755 | -103.6964
16.2792 | -103.7184
16.283 | -103.737
16.2867 | -103.7545
16.2904 | -103.77
16.2941 | -103.7842
16.2978 | -103.8012
16.3016 | -103.8148
16.3053 | -103.8274
16.309 | -103.8458
16.3127 | -103.862
16.3164 | -103.8751
16.3202 | -103.8899
16.3239 | -103.9032
16.3276 | -103.9172
16.3313 | -103.9317
16.335 | -103.9446
16.3388 | -103.9587
16.3425 | -103.9725
16.3462 | -103.9871
16.3499 | -104.0017
16.3536 | -104.0133
16.3573 | -104.0237
16.3611 | -104.0336
16.3648 | -104.0445
16.3685 | -104.0556
16.3722 | -104.0647
16.3759 | -104.0727
16.3797 | -104.0782
16.3834 | -104.0828
16.3871 | -104.0842
16.3908 | -104.0842
16.3945 | -104.0821
16.3983 | -104.0785
16.402 | -104.0758
16.4057 | -104.0739
16.4094 | -104.0694
16.4131 | -104.0596
16.4169 | -104.0479
16.4206 | -104.0363
16.4243 | -104.0255
16.428 | -104.0123
16.4317 | -104.0003
16.4355 | -103.9854
16.4392 | -103.9662
16.4429 | -103.9456
16.4466 | -103.9262
16.4503 | -103.9082
16.454 | -103.891
16.4578 | -103.8678
16.4615 | -103.8453
16.4652 | -103.8264
16.4689 | -103.8085
16.4726 | -103.7886
16.4764 | -103.7676
16.4801 | -103.7517
16.4838 | -103.7389
16.4875 | -103.7257
16.4912 | -103.7157
16.495 | -103.7039
16.4987 | -103.6901
16.5024 | -103.6811
16.5061 | -103.6774
16.5098 | -103.6732
16.5136 | -103.6691
16.5173 | -103.6671
16.521 | -103.6655
16.5247 | -103.6673
16.5284 | -103.6715
16.5322 | -103.6765
16.5359 | -103.6878
16.5396 | -103.6988
16.5433 | -103.7078
16.547 | -103.7197
16.5507 | -103.7309
16.5545 | -103.7396
16.5582 | -103.7505
16.5619 | -103.7631
16.5656 | -103.777
16.5693 | -103.7914
16.5731 | -103.8054
16.5768 | -103.8202
16.5805 | -103.8339
16.5842 | -103.8475
16.5879 | -103.8604
16.5917 | -103.8726
16.5954 | -103.8845
16.5991 | -103.899
16.6028 | -103.9122
16.6065 | -103.9237
16.6103 | -103.9345
16.614 | -103.9463
16.6177 | -103.9594
16.6214 | -103.9734
16.6251 | -103.9868
16.6289 | -104
16.6326 | -104.0127
16.6363 | -104.0259
16.64 | -104.0391
16.6437 | -104.0472
16.6474 | -104.0518
16.6512 | -104.0604
16.6549 | -104.0711
16.6586 | -104.0778
16.6623 | -104.0795
16.666 | -104.0813
16.6698 | -104.0853
16.6735 | -104.0888
16.6772 | -104.091
16.6809 | -104.0928
16.6846 | -104.091
16.6884 | -104.0881
16.6921 | -104.0836
16.6958 | -104.0792
16.6995 | -104.0715
16.7032 | -104.0604
16.707 | -104.0502
16.7107 | -104.04
16.7144 | -104.0268
16.7181 | -104.0122
16.7218 | -103.9979
16.7256 | -103.9822
16.7293 | -103.9666
16.733 | -103.9513
16.7367 | -103.9359
16.7404 | -103.9179
16.7442 | -103.8999
16.7479 | -103.8842
16.7516 | -103.8675
16.7553 | -103.8522
16.759 | -103.8384
16.7627 | -103.8251
16.7665 | -103.8104
16.7702 | -103.7964
16.7739 | -103.7825
16.7776 | -103.772
16.7813 | -103.7631
16.7851 | -103.7526
16.7888 | -103.7458
16.7925 | -103.7389
16.7962 | -103.7336
16.7999 | -103.7289
16.8037 | -103.7261
16.8074 | -103.723
16.8111 | -103.7202
16.8148 | -103.7203
16.8185 | -103.7202
16.8223 | -103.7211
16.826 | -103.7265
16.8297 | -103.7321
16.8334 | -103.7386
16.8371 | -103.7445
16.8409 | -103.7505
16.8446 | -103.7596
16.8483 | -103.7697
16.852 | -103.7804
16.8557 | -103.7897
16.8594 | -103.8011
16.8632 | -103.8153
16.8669 | -103.8316
16.8706 | -103.8486
16.8743 | -103.8656
16.878 | -103.8832
16.8818 | -103.9009
16.8855 | -103.9234
16.8892 | -103.9474
16.8929 | -103.9701
16.8966 | -103.9923
16.9004 | -104.0138
16.9041 | -104.0356
16.9078 | -104.0556
16.9115 | -104.0704
16.9152 | -104.0868
16.919 | -104.1045
16.9227 | -104.1191
16.9264 | -104.1364
16.9301 | -104.1532
16.9338 | -104.169
16.9376 | -104.1822
16.9413 | -104.1955
16.945 | -104.2061
16.9487 | -104.2121
16.9524 | -104.2177
16.9562 | -104.2232
16.9599 | -104.2256
16.9636 | -104.2263
16.9673 | -104.2259
16.971 | -104.2236
16.9747 | -104.2224
16.9785 | -104.221
16.9822 | -104.2177
16.9859 | -104.2118
16.9896 | -104.204
16.9933 | -104.196
16.9971 | -104.1853
17.0008 | -104.1751
17.0045 | -104.1637
17.0082 | -104.155
17.0119 | -104.1433
17.0157 | -104.1307
17.0194 | -104.1187
17.0231 | -104.1054
17.0268 | -104.0929
17.0305 | -104.0809
17.0343 | -104.071
17.038 | -104.0581
17.0417 | -104.0475
17.0454 | -104.0383
17.0491 | -104.0283
17.0529 | -104.0182
17.0566 | -104.0069
17.0603 | -104.0004
17.064 | -103.9924
17.0677 | -103.983
17.0714 | -103.9745
17.0752 | -103.9641
17.0789 | -103.9506
17.0826 | -103.9396
17.0863 | -103.9281
17.09 | -103.9161
17.0938 | -103.9004
17.0975 | -103.8856
17.1012 | -103.8714
17.1049 | -103.8558
17.1086 | -103.8374
17.1124 | -103.8153
17.1161 | -103.7896
17.1198 | -103.7652
17.1235 | -103.7409
17.1272 | -103.7157
17.131 | -103.6909
17.1347 | -103.6638
17.1384 | -103.6363
17.1421 | -103.6101
17.1458 | -103.5839
17.1496 | -103.5585
17.1533 | -103.5327
17.157 | -103.5049
17.1607 | -103.477
17.1644 | -103.4495
17.1681 | -103.4277
17.1719 | -103.402
17.1756 | -103.3796
17.1793 | -103.3589
17.183 | -103.3397
17.1867 | -103.3209
17.1905 | -103.3045
17.1942 | -103.2919
17.1979 | -103.2826
17.2016 | -103.2747
17.2053 | -103.2685
17.2091 | -103.2644
17.2128 | -103.2631
17.2165 | -103.2643
17.2202 | -103.2687
17.2239 | -103.2742
17.2277 | -103.283
17.2314 | -103.2946
17.2351 | -103.3085
17.2388 | -103.3265
17.2425 | -103.3463
17.2463 | -103.3682
17.25 | -103.3931
17.2537 | -103.4172
17.2574 | -103.4441
17.2611 | -103.474
17.2648 | -103.5054
17.2686 | -103.5386
17.2723 | -103.5696
17.276 | -103.6
17.2797 | -103.6322
17.2834 | -103.6657
17.2872 | -103.696
17.2909 | -103.7247
17.2946 | -103.7561
17.2983 | -103.7876
17.302 | -103.8175
17.3058 | -103.8441
17.3095 | -103.8728
17.3132 | -103.8998
17.3169 | -103.9254
17.3206 | -103.9507
17.3244 | -103.9791
17.3281 | -104.0053
17.3318 | -104.0291
17.3355 | -104.0519
17.3392 | -104.0772
17.343 | -104.0998
17.3467 | -104.119
17.3504 | -104.1362
17.3541 | -104.1524
17.3578 | -104.1702
17.3616 | -104.1923
17.3653 | -104.21
17.369 | -104.2265
17.3727 | -104.239
17.3764 | -104.2501
17.3801 | -104.2589
17.3839 | -104.2666
17.3876 | -104.2731
17.3913 | -104.2792
17.395 | -104.2841
17.3987 | -104.2882
17.4025 | -104.2908
17.4062 | -104.2915
17.4099 | -104.2924
17.4136 | -104.2886
17.4173 | -104.2817
17.4211 | -104.2725
17.4248 | -104.2611
17.4285 | -104.2499
17.4322 | -104.2389
17.4359 | -104.2261
17.4397 | -104.2128
17.4434 | -104.1959
17.4471 | -104.1771
17.4508 | -104.155
17.4545 | -104.1312
17.4583 | -104.1076
17.462 | -104.0828
17.4657 | -104.0554
17.4694 | -104.0252
17.4731 | -103.9944
17.4768 | -103.9615
17.4806 | -103.9277
17.4843 | -103.8951
17.488 | -103.8601
17.4917 | -103.8229
17.4954 | -103.7852
17.4992 | -103.7448
17.5029 | -103.7053
17.5066 | -103.6679
17.5103 | -103.6274
17.514 | -103.5851
17.5178 | -103.5424
17.5215 | -103.5032
17.5252 | -103.463
17.5289 | -103.4218
17.5326 | -103.3821
17.5364 | -103.3438
17.5401 | -103.3059
17.5438 | -103.2689
17.5475 | -103.2337
17.5512 | -103.2012
17.555 | -103.1702
17.5587 | -103.1387
17.5624 | -103.1098
17.5661 | -103.0833
17.5698 | -103.0588
17.5735 | -103.0367
17.5773 | -103.0166
17.581 | -103.0012
17.5847 | -102.9893
17.5884 | -102.9777
17.5921 | -102.9681
17.5959 | -102.9619
17.5996 | -102.9566
17.6033 | -102.9539
17.607 | -102.9538
17.6107 | -102.9539
17.6145 | -102.9537
17.6182 | -102.9502
17.6219 | -102.9488
17.6256 | -102.9487
17.6293 | -102.9491
17.6331 | -102.9484
17.6368 | -102.947
17.6405 | -102.944
17.6442 | -102.9397
17.6479 | -102.9358
17.6517 | -102.9305
17.6554 | -102.9219
17.6591 | -102.9097
17.6628 | -102.8972
17.6665 | -102.8842
17.6702 | -102.8701
17.674 | -102.8505
17.6777 | -102.8293
17.6814 | -102.809
17.6851 | -102.7894
17.6888 | -102.769
17.6926 | -102.7486
17.6963 | -102.7283
17.7 | -102.7105
17.7037 | -102.6987
17.7074 | -102.6834
17.7112 | -102.6699
17.7149 | -102.6588
17.7186 | -102.6496
17.7223 | -102.6422
17.726 | -102.6375
17.7298 | -102.6337
17.7335 | -102.6318
17.7372 | -102.6319
17.7409 | -102.6344
17.7446 | -102.6394
17.7484 | -102.6465
17.7521 | -102.6561
17.7558 | -102.6679
17.7595 | -102.6795
17.7632 | -102.6929
17.767 | -102.7085
17.7707 | -102.727
17.7744 | -102.7444
17.7781 | -102.7623
17.7818 | -102.7826
17.7855 | -102.8028
17.7893 | -102.8231
17.793 | -102.841
17.7967 | -102.8609
17.8004 | -102.8802
17.8041 | -102.8977
17.8079 | -102.917
17.8116 | -102.936
17.8153 | -102.9525
17.819 | -102.9677
17.8227 | -102.9793
17.8265 | -102.9872
17.8302 | -102.9956
17.8339 | -103.0019
17.8376 | -103.0073
17.8413 | -103.0144
17.8451 | -103.0199
17.8488 | -103.0235
17.8525 | -103.0276
17.8562 | -103.03
17.8599 | -103.0328
17.8637 | -103.0339
17.8674 | -103.0364
17.8711 | -103.0413
17.8748 | -103.0458
17.8785 | -103.05
17.8822 | -103.0532
17.886 | -103.0561
17.8897 | -103.0596
17.8934 | -103.063
17.8971 | -103.0691
17.9008 | -103.0754
17.9046 | -103.0823
17.9083 | -103.0902
17.912 | -103.1002
17.9157 | -103.1123
17.9194 | -103.1232
17.9232 | -103.1352
17.9269 | -103.1495
17.9306 | -103.1623
17.9343 | -103.1736
17.938 | -103.1865
17.9418 | -103.1985
17.9455 | -103.2121
17.9492 | -103.2256
17.9529 | -103.2385
17.9566 | -103.2528
17.9604 | -103.2664
17.9641 | -103.2795
17.9678 | -103.2896
17.9715 | -103.2984
17.9752 | -103.3084
17.9789 | -103.318
17.9827 | -103.3279
17.9864 | -103.3365
17.9901 | -103.3436
17.9938 | -103.3482
17.9975 | -103.3543
18.0013 | -103.3606
18.005 | -103.3646
18.0087 | -103.3695
18.0124 | -103.3717
18.0161 | -103.3734
18.0199 | -103.3772
18.0236 | -103.3803
18.0273 | -103.3795
18.031 | -103.3785
18.0347 | -103.3791
18.0385 | -103.377
18.0422 | -103.3728
18.0459 | -103.3696
18.0496 | -103.3695
18.0533 | -103.3686
18.0571 | -103.3677
18.0608 | -103.3675
18.0645 | -103.3668
18.0682 | -103.3667
18.0719 | -103.3685
18.0756 | -103.3708
18.0794 | -103.375
18.0831 | -103.378
18.0868 | -103.3828
18.0905 | -103.3878
18.0942 | -103.3945
18.098 | -103.4017
18.1017 | -103.4117
18.1054 | -103.4235
18.1091 | -103.4349
18.1128 | -103.4481
18.1166 | -103.4588
18.1203 | -103.4716
18.124 | -103.4855
18.1277 | -103.5014
18.1314 | -103.5185
18.1352 | -103.5362
18.1389 | -103.5551
18.1426 | -103.573
18.1463 | -103.5918
18.15 | -103.6147
18.1538 | -103.6365
18.1575 | -103.6572
18.1612 | -103.6791
18.1649 | -103.7034
18.1686 | -103.7254
18.1724 | -103.7443
18.1761 | -103.765
18.1798 | -103.7847
18.1835 | -103.8043
18.1872 | -103.8211
18.1909 | -103.8361
18.1947 | -103.8484
18.1984 | -103.8607
18.2021 | -103.871
18.2058 | -103.878
18.2095 | -103.8839
18.2133 | -103.887
18.217 | -103.8896
18.2207 | -103.8891
18.2244 | -103.8868
18.2281 | -103.8816
18.2319 | -103.8724
18.2356 | -103.861
18.2393 | -103.8489
18.243 | -103.8265
18.2467 | -103.8118
18.2505 | -103.7947
18.2542 | -103.7781
18.2579 | -103.7593
18.2616 | -103.7411
18.2653 | -103.7245
18.2691 | -103.7063
18.2728 | -103.6908
18.2765 | -103.6744
18.2802 | -103.6564
18.2839 | -103.6409
18.2876 | -103.6282
18.2914 | -103.6183
18.2951 | -103.6087
18.2988 | -103.6015
18.3025 | -103.5972
18.3062 | -103.595
18.31 | -103.5932
18.3137 | -103.5909
18.3174 | -103.5926
18.3211 | -103.598
18.3248 | -103.6054
18.3286 | -103.6172
18.3323 | -103.6304
18.336 | -103.6451
18.3397 | -103.663
18.3434 | -103.6825
18.3472 | -103.7015
18.3509 | -103.7235
18.3546 | -103.7464
18.3583 | -103.7671
18.362 | -103.7899
18.3658 | -103.8152
18.3695 | -103.8402
18.3732 | -103.868
18.3769 | -103.8976
18.3806 | -103.9262
18.3843 | -103.956
18.3881 | -103.9835
18.3918 | -104.0095
18.3955 | -104.0331
18.3992 | -104.0563
18.4029 | -104.0817
18.4067 | -104.1045
18.4104 | -104.1246
18.4141 | -104.1452
18.4178 | -104.165
18.4215 | -104.1818
18.4253 | -104.1983
18.429 | -104.2157
18.4327 | -104.2314
18.4364 | -104.2449
18.4401 | -104.2567
18.4439 | -104.2685
18.4476 | -104.2823
18.4513 | -104.2932
18.455 | -104.3022
18.4587 | -104.3078
18.4625 | -104.3103
18.4662 | -104.3127
18.4699 | -104.313
18.4736 | -104.313
18.4773 | -104.3119
18.481 | -104.3118
18.4848 | -104.3111
18.4885 | -104.3088
18.4922 | -104.3031
18.4959 | -104.2971
18.4996 | -104.2875
18.5034 | -104.2782
18.5071 | -104.2715
18.5108 | -104.2663
18.5145 | -104.2603
18.5182 | -104.2518
18.522 | -104.241
18.5257 | -104.2299
18.5294 | -104.2165
18.5331 | -104.2049
18.5368 | -104.1909
18.5406 | -104.1794
18.5443 | -104.1698
18.548 | -104.1554
18.5517 | -104.1422
18.5554 | -104.1297
18.5592 | -104.1163
18.5629 | -104.1021
18.5666 | -104.0899
18.5703 | -104.0784
18.574 | -104.0675
18.5778 | -104.057
18.5815 | -104.0455
18.5852 | -104.0333
18.5889 | -104.0202
18.5926 | -104.0063
18.5963 | -103.9961
18.6001 | -103.9868
18.6038 | -103.977
18.6075 | -103.97
18.6112 | -103.9616
18.6149 | -103.9528
18.6187 | -103.946
18.6224 | -103.9418
18.6261 | -103.9375
18.6298 | -103.9344
18.6335 | -103.9339
18.6373 | -103.9313
18.641 | -103.9299
18.6447 | -103.9311
18.6484 | -103.9331
18.6521 | -103.9373
18.6559 | -103.9432
18.6596 | -103.9529
18.6633 | -103.9606
18.667 | -103.9679
18.6707 | -103.977
18.6745 | -103.9855
18.6782 | -103.9938
18.6819 | -104.0045
18.6856 | -104.0164
18.6893 | -104.0283
18.693 | -104.0399
18.6968 | -104.0523
18.7005 | -104.0632
18.7042 | -104.0733
18.7079 | -104.0822
18.7116 | -104.0914
18.7154 | -104.1026
18.7191 | -104.1147
18.7228 | -104.1269
18.7265 | -104.1377
18.7302 | -104.1456
18.734 | -104.1526
18.7377 | -104.1574
18.7414 | -104.1616
18.7451 | -104.1644
18.7488 | -104.1655
18.7526 | -104.1664
18.7563 | -104.1663
18.76 | -104.1643
18.7637 | -104.1613
18.7674 | -104.1577
18.7712 | -104.1521
18.7749 | -104.1455
18.7786 | -104.1318
18.7823 | -104.1219
18.786 | -104.1115
18.7898 | -104.1022
18.7935 | -104.0922
18.7972 | -104.0812
18.8009 | -104.069
18.8046 | -104.0579
18.8083 | -104.0452
18.8121 | -104.0313
18.8158 | -104.0194
18.8195 | -104.0079
18.8232 | -103.9957
18.8269 | -103.986
18.8307 | -103.9775
18.8344 | -103.9668
18.8381 | -103.9552
18.8418 | -103.9456
18.8455 | -103.9393
18.8493 | -103.9346
18.853 | -103.9316
18.8567 | -103.9301
18.8604 | -103.928
18.8641 | -103.9253
18.8679 | -103.9251
18.8716 | -103.9261
18.8753 | -103.9276
18.879 | -103.9302
18.8827 | -103.9344
18.8865 | -103.9403
18.8902 | -103.9485
18.8939 | -103.9572
18.8976 | -103.966
18.9013 | -103.9758
18.905 | -103.9872
18.9088 | -103.9995
18.9125 | -104.0137
18.9162 | -104.0296
18.9199 | -104.0471
18.9236 | -104.0623
18.9274 | -104.0791
18.9311 | -104.1021
18.9348 | -104.1241
18.9385 | -104.145
18.9422 | -104.1674
18.946 | -104.1912
18.9497 | -104.2149
18.9534 | -104.24
18.9571 | -104.263
18.9608 | -104.2862
18.9646 | -104.3099
18.9683 | -104.333
18.972 | -104.3553
18.9757 | -104.3782
18.9794 | -104.4003
18.9832 | -104.4171
18.9869 | -104.4331
18.9906 | -104.4518
18.9943 | -104.4694
18.998 | -104.4864
19.0017 | -104.5009
19.0055 | -104.5127
19.0092 | -104.5223
19.0129 | -104.5307
19.0166 | -104.5374
19.0203 | -104.5406
19.0241 | -104.5436
19.0278 | -104.545
19.0315 | -104.5437
19.0352 | -104.5391
19.0389 | -104.5315
19.0427 | -104.5233
19.0464 | -104.5138
19.0501 | -104.5049
19.0538 | -104.4945
19.0575 | -104.4821
19.0613 | -104.4675
19.065 | -104.4513
19.0687 | -104.4341
19.0724 | -104.4135
19.0761 | -104.3927
19.0799 | -104.3721
19.0836 | -104.3512
19.0873 | -104.3319
19.091 | -104.3093
19.0947 | -104.2865
19.0984 | -104.2635
19.1022 | -104.2421
19.1059 | -104.2243
19.1096 | -104.2073
19.1133 | -104.1894
19.117 | -104.169
19.1208 | -104.1479
19.1245 | -104.129
19.1282 | -104.112
19.1319 | -104.0986
19.1356 | -104.0858
19.1394 | -104.0733
19.1431 | -104.0644
19.1468 | -104.0567
19.1505 | -104.0502
19.1542 | -104.0443
19.158 | -104.0344
19.1617 | -104.0266
19.1654 | -104.022
19.1691 | -104.0163
19.1728 | -104.0122
19.1766 | -104.008
19.1803 | -104.0043
19.184 | -103.9992
19.1877 | -103.9934
19.1914 | -103.9875
19.1952 | -103.9862
19.1989 | -103.9856
19.2026 | -103.9826
19.2063 | -103.978
19.21 | -103.9751
19.2137 | -103.972
19.2175 | -103.9689
19.2212 | -103.9664
19.2249 | -103.9621
19.2286 | -103.9588
19.2323 | -103.9571
19.2361 | -103.9561
19.2398 | -103.9535
19.2435 | -103.9498
19.2472 | -103.9459
19.2509 | -103.9436
19.2547 | -103.9446
19.2584 | -103.9422
19.2621 | -103.9358
19.2658 | -103.9321
19.2695 | -103.9292
19.2733 | -103.9253
19.277 | -103.9214
19.2807 | -103.9172
19.2844 | -103.9151
19.2881 | -103.9141
19.2919 | -103.9139
19.2956 | -103.9131
19.2993 | -103.9098
19.303 | -103.9087
19.3067 | -103.9088
19.3104 | -103.9084
19.3142 | -103.9165
19.3179 | -103.9185
19.3216 | -103.9236
19.3253 | -103.9288
19.329 | -103.9347
19.3328 | -103.9418
19.3365 | -103.9468
19.3402 | -103.9532
19.3439 | -103.9621
19.3476 | -103.9726
19.3514 | -103.984
19.3551 | -103.9946
19.3588 | -104.0052
19.3625 | -104.0147
19.3662 | -104.0272
19.37 | -104.0398
19.3737 | -104.0515
19.3774 | -104.062
19.3811 | -104.0713
19.3848 | -104.0801
19.3886 | -104.0908
19.3923 | -104.1028
19.396 | -104.1147
19.3997 | -104.1231
19.4034 | -104.1304
19.4071 | -104.1371
19.4109 | -104.1442
19.4146 | -104.1499
19.4183 | -104.1526
19.422 | -104.1558
19.4257 | -104.1599
19.4295 | -104.1619
19.4332 | -104.168
19.4369 | -104.1716
19.4406 | -104.1723
19.4443 | -104.173
19.4481 | -104.1728
19.4518 | -104.1733
19.4555 | -104.1729
19.4592 | -104.1709
19.4629 | -104.1684
19.4667 | -104.1668
19.4704 | -104.1635
19.4741 | -104.1565
19.4778 | -104.1505
19.4815 | -104.1472
19.4853 | -104.1429
19.489 | -104.1371
19.4927 | -104.1308
19.4964 | -104.1238
19.5001 | -104.1177
19.5038 | -104.1123
19.5076 | -104.1052
19.5113 | -104.0988
19.515 | -104.094
19.5187 | -104.0901
19.5224 | -104.0828
19.5262 | -104.0747
19.5299 | -104.0685
19.5336 | -104.0645
19.5373 | -104.0647
19.541 | -104.0628
19.5448 | -104.0594
19.5485 | -104.0591
19.5522 | -104.0597
19.5559 | -104.0599
19.5596 | -104.0586
19.5634 | -104.0582
19.5671 | -104.0601
19.5708 | -104.0627
19.5745 | -104.0681
19.5782 | -104.073
19.582 | -104.0775
19.5857 | -104.0823
19.5894 | -104.0853
19.5931 | -104.0892
19.5968 | -104.0943
19.6005 | -104.0997
19.6043 | -104.1074
19.608 | -104.1139
19.6117 | -104.1196
19.6154 | -104.1296
19.6191 | -104.137
19.6229 | -104.1427
19.6266 | -104.1493
19.6303 | -104.1544
19.634 | -104.1582
19.6377 | -104.1619
19.6415 | -104.1664
19.6452 | -104.168
19.6489 | -104.1708
19.6526 | -104.1735
19.6563 | -104.1771
19.6601 | -104.1792
19.6638 | -104.1823
19.6675 | -104.1805
19.6712 | -104.1784
19.6749 | -104.1761
19.6787 | -104.1739
19.6824 | -104.1713
19.6861 | -104.1665
19.6898 | -104.159
19.6935 | -104.1504
19.6973 | -104.1418
19.701 | -104.1312
19.7047 | -104.1191
19.7084 | -104.1075
19.7121 | -104.094
19.7158 | -104.0793
19.7196 | -104.0654
19.7233 | -104.0522
19.727 | -104.0393
19.7307 | -104.025
19.7344 | -104.0105
19.7382 | -103.998
19.7419 | -103.9864
19.7456 | -103.9737
19.7493 | -103.9619
19.753 | -103.9536
19.7568 | -103.946
19.7605 | -103.9386
19.7642 | -103.933
19.7679 | -103.929
19.7716 | -103.9265
19.7754 | -103.9259
19.7791 | -103.9275
19.7828 | -103.9293
19.7865 | -103.934
19.7902 | -103.94
19.794 | -103.9478
19.7977 | -103.9568
19.8014 | -103.9682
19.8051 | -103.9802
19.8088 | -103.9917
19.8125 | -104.0034
19.8163 | -104.0154
19.82 | -104.0293
19.8237 | -104.0463
19.8274 | -104.0621
19.8311 | -104.076
19.8349 | -104.0919
19.8386 | -104.1104
19.8423 | -104.1283
19.846 | -104.1444
19.8497 | -104.158
19.8535 | -104.175
19.8572 | -104.1912
19.8609 | -104.2077
19.8646 | -104.2243
19.8683 | -104.2404
19.8721 | -104.2534
19.8758 | -104.265
19.8795 | -104.2775
19.8832 | -104.2895
19.8869 | -104.2997
19.8907 | -104.3112
19.8944 | -104.3223
19.8981 | -104.3311
19.9018 | -104.3418
19.9055 | -104.3522
19.9093 | -104.3598
19.913 | -104.3662
19.9167 | -104.3741
19.9204 | -104.3824
19.9241 | -104.3895
19.9278 | -104.3962
19.9316 | -104.4004
19.9353 | -104.4043
19.939 | -104.407
19.9427 | -104.4128
19.9464 | -104.418
19.9502 | -104.4256
19.9539 | -104.4328
19.9576 | -104.4379
19.9613 | -104.4442
19.965 | -104.4518
19.9688 | -104.4607
19.9725 | -104.4701
19.9762 | -104.4821
19.9799 | -104.4941
19.9836 | -104.5065
19.9874 | -104.5206
19.9911 | -104.534
19.9948 | -104.5478
19.9985 | -104.5639
20.0022 | -104.5804
20.006 | -104.5984
20.0097 | -104.6141
20.0134 | -104.6304
20.0171 | -104.648
20.0208 | -104.6662
20.0245 | -104.6841
20.0283 | -104.7011
20.032 | -104.7189
20.0357 | -104.7335
20.0394 | -104.7483
20.0431 | -104.7628
20.0469 | -104.777
20.0506 | -104.7911
20.0543 | -104.8049
20.058 | -104.8186
20.0617 | -104.8307
20.0655 | -104.839
20.0692 | -104.8464
20.0729 | -104.8521
20.0766 | -104.8574
20.0803 | -104.8614
20.0841 | -104.8637
20.0878 | -104.8644
20.0915 | -104.8641
20.0952 | -104.8633
20.0989 | -104.8584
20.1027 | -104.8529
20.1064 | -104.8452
20.1101 | -104.8369
20.1138 | -104.8304
20.1175 | -104.8198
20.1212 | -104.8059
20.125 | -104.7899
20.1287 | -104.7745
20.1324 | -104.7597
20.1361 | -104.7455
20.1398 | -104.7272
20.1436 | -104.7092
20.1473 | -104.6924
20.151 | -104.6736
20.1547 | -104.6568
20.1584 | -104.6445
20.1622 | -104.6314
20.1659 | -104.6141
20.1696 | -104.5969
20.1733 | -104.5823
20.177 | -104.5696
20.1808 | -104.5573
20.1845 | -104.5458
20.1882 | -104.535
20.1919 | -104.5254
20.1956 | -104.5156
20.1994 | -104.5054
20.2031 | -104.496
20.2068 | -104.4866
20.2105 | -104.4775
20.2142 | -104.4718
20.2179 | -104.4678
20.2217 | -104.4633
20.2254 | -104.459
20.2291 | -104.4549
20.2328 | -104.4529
20.2365 | -104.4523
20.2403 | -104.4497
20.244 | -104.451
20.2477 | -104.4515
20.2514 | -104.4494
20.2551 | -104.448
20.2589 | -104.4467
20.2626 | -104.4446
20.2663 | -104.4429
20.27 | -104.4422
20.2737 | -104.4415
20.2775 | -104.441
20.2812 | -104.4372
20.2849 | -104.43
20.2886 | -104.424
20.2923 | -104.4185
20.2961 | -104.4144
20.2998 | -104.4106
20.3035 | -104.4031
20.3072 | -104.3943
20.3109 | -104.3863
20.3147 | -104.3776
20.3184 | -104.367
20.3221 | -104.355
20.3258 | -104.3422
20.3295 | -104.3295
20.3332 | -104.3159
20.337 | -104.3008
20.3407 | -104.285
20.3444 | -104.2676
20.3481 | -104.2499
20.3518 | -104.2347
20.3556 | -104.2179
20.3593 | -104.2007
20.363 | -104.1813
20.3667 | -104.1621
20.3704 | -104.1446
20.3742 | -104.1255
20.3779 | -104.1081
20.3816 | -104.0908
20.3853 | -104.0739
20.389 | -104.0567
20.3928 | -104.041
20.3965 | -104.0253
20.4002 | -104.009
20.4039 | -103.9927
20.4076 | -103.9782
20.4114 | -103.9638
20.4151 | -103.9498
20.4188 | -103.9357
20.4225 | -103.9209
20.4262 | -103.9083
20.4299 | -103.8995
20.4337 | -103.8914
20.4374 | -103.8832
20.4411 | -103.8783
20.4448 | -103.8728
20.4485 | -103.8683
20.4523 | -103.8648
20.456 | -103.8646
20.4597 | -103.8669
20.4634 | -103.8698
20.4671 | -103.8722
20.4709 | -103.8754
20.4746 | -103.879
20.4783 | -103.8853
20.482 | -103.8953
20.4857 | -103.9085
20.4895 | -103.9226
20.4932 | -103.9366
20.4969 | -103.9501
20.5006 | -103.965
20.5043 | -103.9804
20.5081 | -103.9951
20.5118 | -104.0126
20.5155 | -104.0296
20.5192 | -104.0434
20.5229 | -104.0555
20.5266 | -104.0679
20.5304 | -104.0826
20.5341 | -104.0954
20.5378 | -104.1071
20.5415 | -104.1186
20.5452 | -104.1277
20.549 | -104.1362
20.5527 | -104.1458
20.5564 | -104.1531
20.5601 | -104.1586
20.5638 | -104.1624
20.5676 | -104.1662
20.5713 | -104.1707
20.575 | -104.1747
20.5787 | -104.1749
20.5824 | -104.1712
20.5862 | -104.1671
20.5899 | -104.1632
20.5936 | -104.1607
20.5973 | -104.1566
20.601 | -104.1498
20.6048 | -104.1441
20.6085 | -104.1401
20.6122 | -104.1382
20.6159 | -104.1367
20.6196 | -104.1352
20.6233 | -104.1329
20.6271 | -104.1319
20.6308 | -104.1319
20.6345 | -104.1325
20.6382 | -104.1338
20.6419 | -104.1347
20.6457 | -104.1371
20.6494 | -104.1409
20.6531 | -104.1479
20.6568 | -104.1559
20.6605 | -104.1632
20.6643 | -104.1697
20.668 | -104.1754
20.6717 | -104.1832
20.6754 | -104.191
20.6791 | -104.1997
20.6829 | -104.2096
20.6866 | -104.2192
20.6903 | -104.2271
20.694 | -104.2367
20.6977 | -104.2474
20.7015 | -104.2573
20.7052 | -104.2686
20.7089 | -104.281
20.7126 | -104.2932
20.7163 | -104.3052
20.7201 | -104.3133
20.7238 | -104.3203
20.7275 | -104.3279
20.7312 | -104.3337
20.7349 | -104.3389
20.7386 | -104.3457
20.7424 | -104.3512
20.7461 | -104.3555
20.7498 | -104.3584
20.7535 | -104.359
20.7572 | -104.36
20.761 | -104.36
20.7647 | -104.362
20.7684 | -104.3637
20.7721 | -104.3609
20.7758 | -104.3548
20.7796 | -104.3486
20.7833 | -104.343
20.787 | -104.3374
20.7907 | -104.3316
20.7944 | -104.3259
20.7982 | -104.3177
20.8019 | -104.3114
20.8056 | -104.3039
20.8093 | -104.2976
20.813 | -104.2936
20.8168 | -104.2911
20.8205 | -104.2865
20.8242 | -104.2807
20.8279 | -104.2753
20.8316 | -104.2719
20.8353 | -104.2693
20.8391 | -104.2684
20.8428 | -104.2683
20.8465 | -104.266
20.8502 | -104.2668
20.8539 | -104.2653
20.8577 | -104.2647
20.8614 | -104.2665
20.8651 | -104.2707
20.8688 | -104.276
20.8725 | -104.2822
20.8763 | -104.2901
20.88 | -104.2997
20.8837 | -104.3084
20.8874 | -104.318
20.8911 | -104.3284
20.8949 | -104.3417
20.8986 | -104.3541
20.9023 | -104.3639
20.906 | -104.3747
20.9097 | -104.3837
20.9135 | -104.3925
20.9172 | -104.4021
20.9209 | -104.4128
20.9246 | -104.4202
20.9283 | -104.4308
20.932 | -104.4415
20.9358 | -104.4507
20.9395 | -104.4597
20.9432 | -104.4678
20.9469 | -104.474
20.9506 | -104.4782
20.9544 | -104.4812
20.9581 | -104.4855
20.9618 | -104.4935
20.9655 | -104.4991
20.9692 | -104.5029
20.973 | -104.5061
20.9767 | -104.5079
20.9804 | -104.5088
20.9841 | -104.5074
20.9878 | -104.5059
20.9916 | -104.5047
20.9953 | -104.502
20.999 | -104.5016
21.0027 | -104.5002
21.0064 | -104.4979
21.0102 | -104.4952
21.0139 | -104.4913
21.0176 | -104.4891
21.0213 | -104.4865
21.025 | -104.4842
21.0287 | -104.4826
21.0325 | -104.482
21.0362 | -104.4806
21.0399 | -104.4769
21.0436 | -104.4749
21.0473 | -104.4741
21.0511 | -104.4739
21.0548 | -104.4734
21.0585 | -104.4729
21.0622 | -104.472
21.0659 | -104.4718
21.0697 | -104.4734
21.0734 | -104.4765
21.0771 | -104.4793
21.0808 | -104.4821
21.0845 | -104.4841
21.0883 | -104.4866
21.092 | -104.4886
21.0957 | -104.4914
21.0994 | -104.4928
21.1031 | -104.4935
21.1069 | -104.4946
21.1106 | -104.4965
21.1143 | -104.4957
21.118 | -104.4938
21.1217 | -104.4924
21.1255 | -104.491
21.1292 | -104.491
21.1329 | -104.4914
21.1366 | -104.4905
21.1403 | -104.4893
21.144 | -104.4887
21.1478 | -104.4856
21.1515 | -104.4843
21.1552 | -104.4843
21.1589 | -104.4822
21.1626 | -104.4796
21.1664 | -104.4785
21.1701 | -104.4777
21.1738 | -104.4749
21.1775 | -104.4727
21.1812 | -104.4702
21.185 | -104.4685
21.1887 | -104.4685
21.1924 | -104.4674
21.1961 | -104.4683
21.1998 | -104.4709
21.2036 | -104.4701
21.2073 | -104.4711
21.211 | -104.4738
21.2147 | -104.4751
21.2184 | -104.4763
21.2222 | -104.4772
21.2259 | -104.4767
21.2296 | -104.4783
21.2333 | -104.4811
21.237 | -104.4826
21.2407 | -104.4849
21.2445 | -104.4882
21.2482 | -104.4913
21.2519 | -104.4959
21.2556 | -104.5018
21.2593 | -104.5068
21.2631 | -104.5102
21.2668 | -104.5165
21.2705 | -104.5234
21.2742 | -104.5316
21.2779 | -104.5396
21.2817 | -104.546
21.2854 | -104.5545
21.2891 | -104.5643
21.2928 | -104.572
21.2965 | -104.5795
21.3003 | -104.5866
21.304 | -104.5938
21.3077 | -104.602
21.3114 | -104.6108
21.3151 | -104.6199
21.3189 | -104.6291
21.3226 | -104.6379
21.3263 | -104.6474
21.33 | -104.6575
21.3337 | -104.6655
21.3374 | -104.6745
21.3412 | -104.6865
21.3449 | -104.6992
21.3486 | -104.7096
21.3523 | -104.7199
21.356 | -104.7296
21.3598 | -104.7416
21.3635 | -104.7514
21.3672 | -104.761
21.3709 | -104.7699
21.3746 | -104.779
21.3784 | -104.7903
21.3821 | -104.7998
21.3858 | -104.809
21.3895 | -104.8183
21.3932 | -104.8246
21.397 | -104.8291
21.4007 | -104.8355
21.4044 | -104.8427
21.4081 | -104.8501
21.4118 | -104.8554
21.4156 | -104.8604
21.4193 | -104.8662
21.423 | -104.8726
21.4267 | -104.8763
21.4304 | -104.8767
21.4341 | -104.8755
21.4379 | -104.8716
21.4416 | -104.8687
21.4453 | -104.8681
21.449 | -104.8688
21.4527 | -104.8692
21.4565 | -104.8646
21.4602 | -104.8503
21.4639 | -104.8404
21.4676 | -104.8309
21.4713 | -104.8201
21.4751 | -104.8097
21.4788 | -104.7994
21.4825 | -104.7881
21.4862 | -104.7774
21.4899 | -104.7648
21.4937 | -104.7497
21.4974 | -104.7342
21.5011 | -104.7185
21.5048 | -104.7046
21.5085 | -104.6931
21.5123 | -104.6792
21.516 | -104.6638
21.5197 | -104.6481
21.5234 | -104.6325
21.5271 | -104.6171
21.5309 | -104.6034
21.5346 | -104.5881
21.5383 | -104.5735
21.542 | -104.5611
21.5457 | -104.5457
21.5494 | -104.5321
21.5532 | -104.518
21.5569 | -104.5008
21.5606 | -104.4845
21.5643 | -104.4702
21.568 | -104.4531
21.5718 | -104.4376
21.5755 | -104.4229
21.5792 | -104.411
21.5829 | -104.3982
21.5866 | -104.3867
21.5904 | -104.3765
21.5941 | -104.3665
21.5978 | -104.3568
21.6015 | -104.3476
21.6052 | -104.3376
21.609 | -104.3288
21.6127 | -104.3201
21.6164 | -104.3113
21.6201 | -104.3054
21.6238 | -104.2996
21.6276 | -104.2934
21.6313 | -104.2867
21.635 | -104.2813
21.6387 | -104.2769
21.6424 | -104.2719
21.6461 | -104.2686
21.6499 | -104.2642
21.6536 | -104.2593
21.6573 | -104.2554
21.661 | -104.2517
21.6647 | -104.2476
21.6685 | -104.2466
21.6722 | -104.2438
21.6759 | -104.2399
21.6796 | -104.2367
21.6833 | -104.2337
21.6871 | -104.2337
21.6908 | -104.2349
21.6945 | -104.2353
21.6982 | -104.2362
21.7019 | -104.2366
21.7057 | -104.2361
21.7094 | -104.2346
21.7131 | -104.2315
21.7168 | -104.2287
21.7205 | -104.2279
21.7243 | -104.2282
21.728 | -104.2262
21.7317 | -104.2241
21.7354 | -104.2217
21.7391 | -104.22
21.7429 | -104.2198
21.7466 | -104.2171
21.7503 | -104.2096
21.754 | -104.2024
21.7577 | -104.1976
21.7614 | -104.1946
21.7652 | -104.1918
21.7689 | -104.1894
21.7726 | -104.1856
21.7763 | -104.1811
21.78 | -104.1767
21.7838 | -104.1697
21.7875 | -104.1641
21.7912 | -104.1586
21.7949 | -104.1529
21.7986 | -104.1462
21.8024 | -104.1402
21.8061 | -104.1366
21.8098 | -104.1312
21.8135 | -104.1251
21.8172 | -104.1197
21.821 | -104.1182
21.8247 | -104.1158
21.8284 | -104.1135
21.8321 | -104.1111
21.8358 | -104.1101
21.8396 | -104.1086
21.8433 | -104.1067
21.847 | -104.1044
21.8507 | -104.1048
21.8544 | -104.1078
21.8581 | -104.1099
21.8619 | -104.1142
21.8656 | -104.1213
21.8693 | -104.1267
21.873 | -104.1321
21.8767 | -104.141
21.8805 | -104.1499
21.8842 | -104.1588
21.8879 | -104.1669
21.8916 | -104.1767
21.8953 | -104.187
21.8991 | -104.1973
21.9028 | -104.2085
21.9065 | -104.2207
21.9102 | -104.2326
21.9139 | -104.2444
21.9177 | -104.2572
21.9214 | -104.2721
21.9251 | -104.2864
21.9288 | -104.2974
21.9325 | -104.3099
21.9363 | -104.3241
21.94 | -104.339
21.9437 | -104.3532
21.9474 | -104.3667
21.9511 | -104.3801
21.9548 | -104.3918
21.9586 | -104.4039
21.9623 | -104.4159
21.966 | -104.4273
21.9697 | -104.4403
21.9734 | -104.4512
21.9772 | -104.4575
21.9809 | -104.4628
21.9846 | -104.4689
21.9883 | -104.4735
21.992 | -104.4762
21.9958 | -104.4773
21.9995 | -104.4725
22.0032 | -104.4723
22.0069 | -104.4715
22.0106 | -104.4709
22.0144 | -104.4686
22.0181 | -104.4632
22.0218 | -104.4583
22.0255 | -104.4535
22.0292 | -104.4468
22.033 | -104.4384
22.0367 | -104.4294
22.0404 | -104.4202
22.0441 | -104.4115
22.0478 | -104.4005
22.0515 | -104.3903
22.0553 | -104.3808
22.059 | -104.3709
22.0627 | -104.3629
22.0664 | -104.3535
22.0701 | -104.3418
22.0739 | -104.3298
22.0776 | -104.3181
22.0813 | -104.3051
22.085 | -104.2959
22.0887 | -104.2862
22.0925 | -104.2763
22.0962 | -104.2688
22.0999 | -104.2618
22.1036 | -104.2542
22.1073 | -104.2468
22.1111 | -104.2422
22.1148 | -104.2374
22.1185 | -104.2319
22.1222 | -104.2281
22.1259 | -104.2249
22.1297 | -104.2235
22.1334 | -104.2208
22.1371 | -104.2215
22.1408 | -104.2226
22.1445 | -104.2241
22.1483 | -104.2267
22.152 | -104.2321
22.1557 | -104.2375
22.1594 | -104.2434
22.1631 | -104.2501
22.1668 | -104.2591
22.1706 | -104.2667
22.1743 | -104.2748
22.178 | -104.2843
22.1817 | -104.2944
22.1854 | -104.3044
22.1892 | -104.3159
22.1929 | -104.3244
22.1966 | -104.3333
22.2003 | -104.3442
22.204 | -104.3563
22.2078 | -104.3693
22.2115 | -104.382
22.2152 | -104.3945
22.2189 | -104.4062
22.2226 | -104.4178
22.2264 | -104.4306
22.2301 | -104.4428
22.2338 | -104.4536
22.2375 | -104.4661
22.2412 | -104.4794
22.245 | -104.4908
22.2487 | -104.5032
22.2524 | -104.5143
22.2561 | -104.5254
22.2598 | -104.5379
22.2635 | -104.5501
22.2673 | -104.5622
22.271 | -104.5738
22.2747 | -104.5852
22.2784 | -104.5956
22.2821 | -104.6054
22.2859 | -104.6177
22.2896 | -104.6314
22.2933 | -104.6464
22.297 | -104.6596
22.3007 | -104.6711
22.3045 | -104.6831
22.3082 | -104.6938
22.3119 | -104.7038
22.3156 | -104.7134
22.3193 | -104.7246
22.3231 | -104.7363
22.3268 | -104.7492
22.3305 | -104.7634
22.3342 | -104.7771
22.3379 | -104.7877
22.3417 | -104.7963
22.3454 | -104.8054
22.3491 | -104.8152
22.3528 | -104.825
22.3565 | -104.8344
22.3602 | -104.8424
22.364 | -104.8518
22.3677 | -104.8592
22.3714 | -104.8662
22.3751 | -104.873
22.3788 | -104.8802
22.3826 | -104.8873
22.3863 | -104.8924
22.39 | -104.898
22.3937 | -104.9041
22.3974 | -104.9124
22.4012 | -104.918
22.4049 | -104.9203
22.4086 | -104.9214
22.4123 | -104.9207
22.416 | -104.9203
22.4198 | -104.9214
22.4235 | -104.9217
22.4272 | -104.9212
22.4309 | -104.9205
22.4346 | -104.9172
22.4384 | -104.9125
22.4421 | -104.908
22.4458 | -104.9024
22.4495 | -104.8974
22.4532 | -104.8901
22.4569 | -104.886
22.4607 | -104.8808
22.4644 | -104.8735
22.4681 | -104.8655
22.4718 | -104.8593
22.4755 | -104.8535
22.4793 | -104.8461
22.483 | -104.8364
22.4867 | -104.8285
22.4904 | -104.8217
22.4941 | -104.8147
22.4979 | -104.8075
22.5016 | -104.8006
22.5053 | -104.7934
22.509 | -104.7894
22.5127 | -104.7862
22.5165 | -104.7807
22.5202 | -104.7765
22.5239 | -104.7714
22.5276 | -104.7663
22.5313 | -104.7633
22.5351 | -104.7686
22.5388 | -104.7683
22.5425 | -104.7695
22.5462 | -104.7701
22.5499 | -104.7696
22.5537 | -104.768
22.5574 | -104.7672
22.5611 | -104.7693
22.5648 | -104.7719
22.5685 | -104.773
22.5722 | -104.7756
22.576 | -104.7788
22.5797 | -104.7827
22.5834 | -104.7868
22.5871 | -104.791
22.5908 | -104.7948
22.5946 | -104.8
22.5983 | -104.8028
22.602 | -104.8036
22.6057 | -104.8032
22.6094 | -104.8051
22.6132 | -104.808
22.6169 | -104.8099
22.6206 | -104.8108
22.6243 | -104.8119
22.628 | -104.8139
22.6318 | -104.816
22.6355 | -104.8179
22.6392 | -104.8175
22.6429 | -104.8183
22.6466 | -104.8197
22.6504 | -104.8201
22.6541 | -104.82
22.6578 | -104.8217
22.6615 | -104.8233
22.6652 | -104.8241
22.6689 | -104.8228
22.6727 | -104.8199
22.6764 | -104.817
22.6801 | -104.8154
22.6838 | -104.8161
22.6875 | -104.8156
22.6913 | -104.8169
22.695 | -104.8178
22.6987 | -104.8179
22.7024 | -104.8194
22.7061 | -104.8228
22.7099 | -104.8246
22.7136 | -104.8264
22.7173 | -104.8276
22.721 | -104.8271
22.7247 | -104.8299
22.7285 | -104.833
22.7322 | -104.8372
22.7359 | -104.8406
22.7396 | -104.8411
22.7433 | -104.8442
22.7471 | -104.8462
22.7508 | -104.8473
22.7545 | -104.8504
22.7582 | -104.8558
22.7619 | -104.8609
22.7656 | -104.8652
22.7694 | -104.8672
22.7731 | -104.8704
22.7768 | -104.8731
22.7805 | -104.8755
22.7842 | -104.8787
22.788 | -104.8806
22.7917 | -104.8821
22.7954 | -104.8842
22.7991 | -104.8881
22.8028 | -104.8911
22.8066 | -104.893
22.8103 | -104.8955
22.814 | -104.8981
22.8177 | -104.8998
22.8214 | -104.9017
22.8252 | -104.9026
22.8289 | -104.9011
22.8326 | -104.8993
22.8363 | -104.8989
22.84 | -104.8972
22.8438 | -104.8946
22.8475 | -104.8932
22.8512 | -104.8917
22.8549 | -104.8876
22.8586 | -104.883
22.8623 | -104.8806
22.8661 | -104.8779
22.8698 | -104.8741
22.8735 | -104.8722
22.8772 | -104.8711
22.8809 | -104.8676
22.8847 | -104.8639
22.8884 | -104.8601
22.8921 | -104.8571
22.8958 | -104.8543
22.8995 | -104.8515
22.9033 | -104.8472
22.907 | -104.843
22.9107 | -104.8387
22.9144 | -104.8353
22.9181 | -104.8315
22.9219 | -104.8278
22.9256 | -104.8262
22.9293 | -104.8265
22.933 | -104.8232
22.9367 | -104.818
22.9405 | -104.8147
22.9442 | -104.8127
22.9479 | -104.8097
22.9516 | -104.8074
22.9553 | -104.8037
22.9591 | -104.8007
22.9628 | -104.7993
22.9665 | -104.7961
22.9702 | -104.7935
22.9739 | -104.7928
22.9776 | -104.7928
22.9814 | -104.7906
22.9851 | -104.7901
22.9888 | -104.7903
22.9925 | -104.7907
22.9962 | -104.7914
23 | -104.7916
23.0037 | -104.7897
23.0074 | -104.7876
23.0111 | -104.7882
23.0148 | -104.7861
23.0186 | -104.7862
23.0223 | -104.7864
23.026 | -104.7848
23.0297 | -104.7831
23.0334 | -104.7827
23.0372 | -104.7829
23.0409 | -104.7842
23.0446 | -104.7867
23.0483 | -104.7875
23.052 | -104.788
23.0558 | -104.7872
23.0595 | -104.786
23.0632 | -104.7863
23.0669 | -104.7828
23.0706 | -104.7792
23.0743 | -104.7783
23.0781 | -104.7783
23.0818 | -104.777
23.0855 | -104.7742
23.0892 | -104.7703
23.0929 | -104.7681
23.0967 | -104.767
23.1004 | -104.7643
23.1041 | -104.7617
23.1078 | -104.7612
23.1115 | -104.7588
23.1153 | -104.7566
23.119 | -104.7526
23.1227 | -104.7486
23.1264 | -104.7439
23.1301 | -104.7407
23.1339 | -104.7365
23.1376 | -104.7317
23.1413 | -104.729
23.145 | -104.7291
23.1487 | -104.7271
23.1525 | -104.727
23.1562 | -104.7287
23.1599 | -104.7286
23.1636 | -104.7256
23.1673 | -104.7248
23.171 | -104.7257
23.1748 | -104.7242
23.1785 | -104.7217
23.1822 | -104.7204
23.1859 | -104.7203
23.1896 | -104.7173
23.1934 | -104.7153
23.1971 | -104.7149
23.2008 | -104.7132
23.2045 | -104.7122
23.2082 | -104.713
23.212 | -104.7115
23.2157 | -104.7098
23.2194 | -104.7066
23.2231 | -104.7033
23.2268 | -104.7005
23.2306 | -104.6989
23.2343 | -104.6989
23.238 | -104.7003
23.2417 | -104.7006
23.2454 | -104.7011
23.2492 | -104.7024
23.2529 | -104.7013
23.2566 | -104.6997
23.2603 | -104.7011
23.264 | -104.7019
23.2677 | -104.7023
23.2715 | -104.7025
23.2752 | -104.7021
23.2789 | -104.7034
23.2826 | -104.7068
23.2863 | -104.7092
23.2901 | -104.711
23.2938 | -104.7135
23.2975 | -104.7163
23.3012 | -104.7184
23.3049 | -104.7197
23.3087 | -104.7223
23.3124 | -104.7259
23.3161 | -104.7287
23.3198 | -104.7332
23.3235 | -104.7366
23.3273 | -104.7398
23.331 | -104.7459
23.3347 | -104.7509
23.3384 | -104.7532
23.3421 | -104.7541
23.3459 | -104.7553
23.3496 | -104.7576
23.3533 | -104.7627
23.357 | -104.768
23.3607 | -104.7721
23.3645 | -104.7756
23.3682 | -104.7783
23.3719 | -104.7809
23.3756 | -104.7847
23.3793 | -104.7884
23.383 | -104.7927
23.3868 | -104.7961
23.3905 | -104.7988
23.3942 | -104.8021
23.3979 | -104.805
23.4016 | -104.8064
23.4054 | -104.8059
23.4091 | -104.807
23.4128 | -104.809
23.4165 | -104.8102
23.4202 | -104.8128
23.424 | -104.8156
23.4277 | -104.8184
23.4314 | -104.8207
23.4351 | -104.8228
23.4388 | -104.8242
23.4426 | -104.8272
23.4463 | -104.8306
23.45 | -104.8364
23.4537 | -104.8413
23.4574 | -104.8451
23.4612 | -104.8487
23.4649 | -104.8563
23.4686 | -104.8647
23.4723 | -104.8693
23.476 | -104.874
23.4797 | -104.8802
23.4835 | -104.8864
23.4872 | -104.8912
23.4909 | -104.8982
23.4946 | -104.9081
23.4983 | -104.9187
23.5021 | -104.9292
23.5058 | -104.94
23.5095 | -104.9514
23.5132 | -104.9623
23.5169 | -104.9721
23.5207 | -104.9845
23.5244 | -104.9968
23.5281 | -105.0091
23.5318 | -105.0217
23.5355 | -105.0338
23.5393 | -105.0448
23.543 | -105.0556
23.5467 | -105.0671
23.5504 | -105.0794
23.5541 | -105.0914
23.5579 | -105.1027
23.5616 | -105.1117
23.5653 | -105.1197
23.569 | -105.1273
23.5727 | -105.1359
23.5765 | -105.1446
23.5802 | -105.1522
23.5839 | -105.1591
23.5876 | -105.1643
23.5913 | -105.1697
23.595 | -105.1763
23.5988 | -105.1815
23.6025 | -105.1872
23.6062 | -105.1889
23.6099 | -105.1941
23.6136 | -105.1973
23.6174 | -105.2002
23.6211 | -105.2033
23.6248 | -105.2063
23.6285 | -105.207
23.6322 | -105.206
23.636 | -105.2088
23.6397 | -105.2127
23.6434 | -105.2161
23.6471 | -105.2171
23.6508 | -105.2183
23.6546 | -105.2177
23.6583 | -105.2185
23.662 | -105.2219
23.6657 | -105.2228
23.6694 | -105.2232
23.6732 | -105.2255
23.6769 | -105.2301
23.6806 | -105.2331
23.6843 | -105.2348
23.688 | -105.2331
23.6917 | -105.2321
23.6955 | -105.2312
23.6992 | -105.2294
23.7029 | -105.2278
23.7066 | -105.2274
23.7103 | -105.2269
23.7141 | -105.2267
23.7178 | -105.2268
23.7215 | -105.2273
23.7252 | -105.2283
23.7289 | -105.2288
23.7327 | -105.2268
23.7364 | -105.2265
23.7401 | -105.2239
23.7438 | -105.2187
23.7475 | -105.2144
23.7513 | -105.21
23.755 | -105.2032
23.7587 | -105.1965
23.7624 | -105.1903
23.7661 | -105.1855
23.7699 | -105.1795
23.7736 | -105.1731
23.7773 | -105.1673
23.781 | -105.1602
23.7847 | -105.1541
23.7884 | -105.1476
23.7922 | -105.1401
23.7959 | -105.1329
23.7996 | -105.1246
23.8033 | -105.1144
23.807 | -105.1063
23.8108 | -105.0976
23.8145 | -105.0899
23.8182 | -105.0827
23.8219 | -105.0775
23.8256 | -105.0721
23.8294 | -105.065
23.8331 | -105.0547
23.8368 | -105.0451
23.8405 | -105.0382
23.8442 | -105.0306
23.848 | -105.0229
23.8517 | -105.0172
23.8554 | -105.0116
23.8591 | -105.0054
23.8628 | -104.9998
23.8666 | -104.9948
23.8703 | -104.9899
23.874 | -104.9865
23.8777 | -104.9846
23.8814 | -104.9822
23.8851 | -104.9792
23.8889 | -104.9773
23.8926 | -104.9772
23.8963 | -104.9778
23.9 | -104.9766
23.9037 | -104.9765
23.9075 | -104.9768
23.9112 | -104.978
23.9149 | -104.9809
23.9186 | -104.9856
23.9223 | -104.9898
23.9261 | -104.9922
23.9298 | -104.9963
23.9335 | -105.0024
23.9372 | -105.0076
23.9409 | -105.0125
23.9447 | -105.0187
23.9484 | -105.0268
23.9521 | -105.0337
23.9558 | -105.0408
23.9595 | -105.0489
23.9633 | -105.0585
23.967 | -105.0696
23.9707 | -105.0807
23.9744 | -105.0893
23.9781 | -105.0988
23.9819 | -105.1084
23.9856 | -105.1188
23.9893 | -105.1294
23.993 | -105.1396
23.9967 | -105.1495
24.0004 | -105.1583
24.0042 | -105.1674
24.0079 | -105.1764
24.0116 | -105.1847
24.0153 | -105.1928
24.019 | -105.2018
24.0228 | -105.2104
24.0265 | -105.2198
24.0302 | -105.2277
24.0339 | -105.2366
24.0376 | -105.2447
24.0414 | -105.2517
24.0451 | -105.2582
24.0488 | -105.262
24.0525 | -105.2649
24.0562 | -105.2676
24.06 | -105.2708
24.0637 | -105.2745
24.0674 | -105.2794
24.0711 | -105.2843
24.0748 | -105.2881
24.0786 | -105.2928
24.0823 | -105.2956
24.086 | -105.2963
24.0897 | -105.2968
24.0934 | -105.2988
24.0971 | -105.2977
24.1009 | -105.2968
24.1046 | -105.2946
24.1083 | -105.2924
24.112 | -105.2906
24.1157 | -105.2872
24.1195 | -105.2836
24.1232 | -105.2777
24.1269 | -105.2731
24.1306 | -105.2689
24.1343 | -105.2644
24.1381 | -105.2598
24.1418 | -105.2554
24.1455 | -105.2487
24.1492 | -105.2448
24.1529 | -105.2403
24.1567 | -105.2355
24.1604 | -105.2284
24.1641 | -105.2212
24.1678 | -105.2134
24.1715 | -105.2076
24.1753 | -105.2044
24.179 | -105.2004
24.1827 | -105.1945
24.1864 | -105.1905
24.1901 | -105.1865
24.1938 | -105.1816
24.1976 | -105.1761
24.2013 | -105.17
24.205 | -105.1632
24.2087 | -105.156
24.2124 | -105.1491
24.2162 | -105.1428
24.2199 | -105.1377
24.2236 | -105.1334
24.2273 | -105.1308
24.231 | -105.1275
24.2348 | -105.1231
24.2385 | -105.1199
24.2422 | -105.118
24.2459 | -105.1144
24.2496 | -105.1122
24.2534 | -105.1083
24.2571 | -105.1054
24.2608 | -105.1042
24.2645 | -105.102
24.2682 | -105.1009
24.272 | -105.0994
24.2757 | -105.0964
24.2794 | -105.0937
24.2831 | -105.0906
24.2868 | -105.087
24.2905 | -105.0814
24.2943 | -105.0769
24.298 | -105.0719
24.3017 | -105.067
24.3054 | -105.0589
24.3091 | -105.0483
24.3129 | -105.0356
24.3166 | -105.0233
24.3203 | -105.0099
24.324 | -104.9976
24.3277 | -104.9858
24.3315 | -104.9728
24.3352 | -104.9601
24.3389 | -104.946
24.3426 | -104.9335
24.3463 | -104.9194
24.3501 | -104.9045
24.3538 | -104.889
24.3575 | -104.8738
24.3612 | -104.8566
24.3649 | -104.8382
24.3687 | -104.8191
24.3724 | -104.7981
24.3761 | -104.7747
24.3798 | -104.7551
24.3835 | -104.7353
24.3873 | -104.7138
24.391 | -104.6937
24.3947 | -104.6736
24.3984 | -104.6525
24.4021 | -104.6309
24.4058 | -104.6092
24.4096 | -104.5882
24.4133 | -104.5672
24.417 | -104.5468
24.4207 | -104.5273
24.4244 | -104.5088
24.4282 | -104.4907
24.4319 | -104.474
24.4356 | -104.4582
24.4393 | -104.4432
24.443 | -104.4286
24.4468 | -104.4148
24.4505 | -104.4021
24.4542 | -104.3902
24.4579 | -104.3807
24.4616 | -104.3726
24.4654 | -104.3652
24.4691 | -104.3596
24.4728 | -104.3553
24.4765 | -104.3523
24.4802 | -104.3506
24.484 | -104.3501
24.4877 | -104.3519
24.4914 | -104.3531
24.4951 | -104.3548
24.4988 | -104.3571
24.5025 | -104.3604
24.5063 | -104.3678
24.51 | -104.3758
24.5137 | -104.3846
24.5174 | -104.393
24.5211 | -104.4017
24.5249 | -104.4139
24.5286 | -104.4268
24.5323 | -104.4395
24.536 | -104.4535
24.5397 | -104.4668
24.5435 | -104.4808
24.5472 | -104.4951
24.5509 | -104.51
24.5546 | -104.5255
24.5583 | -104.5402
24.5621 | -104.5532
24.5658 | -104.5638
24.5695 | -104.5756
24.5732 | -104.5858
24.5769 | -104.5924
24.5807 | -104.5949
24.5844 | -104.599
24.5881 | -104.6057
24.5918 | -104.6118
24.5955 | -104.6153
24.5992 | -104.6194
24.603 | -104.6237
24.6067 | -104.6257
24.6104 | -104.6267
24.6141 | -104.6257
24.6178 | -104.6242
24.6216 | -104.6218
24.6253 | -104.6202
24.629 | -104.6195
24.6327 | -104.6181
24.6364 | -104.616
24.6402 | -104.6135
24.6439 | -104.6089
24.6476 | -104.6029
24.6513 | -104.5997
24.655 | -104.5984
24.6588 | -104.5966
24.6625 | -104.5926
24.6662 | -104.5879
24.6699 | -104.5816
24.6736 | -104.5739
24.6774 | -104.5678
24.6811 | -104.5635
24.6848 | -104.5565
24.6885 | -104.5507
24.6922 | -104.5435
24.6959 | -104.5361
24.6997 | -104.5312
24.7034 | -104.5279
24.7071 | -104.5231
24.7108 | -104.5198
24.7145 | -104.5182
24.7183 | -104.5164
24.722 | -104.5169
24.7257 | -104.5183
24.7294 | -104.5168
24.7331 | -104.515
24.7369 | -104.5137
24.7406 | -104.5123
24.7443 | -104.5138
24.748 | -104.5138
24.7517 | -104.515
24.7555 | -104.5173
24.7592 | -104.5195
24.7629 | -104.5222
24.7666 | -104.5257
24.7703 | -104.5282
24.7741 | -104.53
24.7778 | -104.5343
24.7815 | -104.5386
24.7852 | -104.5421
24.7889 | -104.5469
24.7927 | -104.5525
24.7964 | -104.5595
24.8001 | -104.5661
24.8038 | -104.5725
24.8075 | -104.5772
24.8112 | -104.5815
24.815 | -104.5837
24.8187 | -104.5796
24.8224 | -104.5694
24.8261 | -104.5528
24.8298 | -104.5425
24.8336 | -104.5389
24.8373 | -104.5351
24.841 | -104.5315
24.8447 | -104.5273
24.8484 | -104.5223
24.8522 | -104.5154
24.8559 | -104.5047
24.8596 | -104.4941
24.8633 | -104.4836
24.867 | -104.4721
24.8708 | -104.4583
24.8745 | -104.4421
24.8782 | -104.4262
24.8819 | -104.4095
24.8856 | -104.3908
24.8894 | -104.3719
24.8931 | -104.3533
24.8968 | -104.3345
24.9005 | -104.3138
24.9042 | -104.2923
24.9079 | -104.2712
24.9117 | -104.2491
24.9154 | -104.2283
24.9191 | -104.2053
24.9228 | -104.1839
24.9265 | -104.1659
24.9303 | -104.1494
24.934 | -104.1314
24.9377 | -104.1107
24.9414 | -104.0917
24.9451 | -104.0741
24.9489 | -104.0555
24.9526 | -104.0369
24.9563 | -104.0214
24.96 | -104.0087
24.9637 | -103.9969
24.9675 | -103.9845
24.9712 | -103.976
24.9749 | -103.9686
24.9786 | -103.9631
24.9823 | -103.9584
24.9861 | -103.9558
24.9898 | -103.9547
24.9935 | -103.9574
24.9972 | -103.9629
25.0009 | -103.9688
25.0046 | -103.9767
25.0084 | -103.9865
25.0121 | -103.9981
25.0158 | -104.0125
25.0195 | -104.0288
25.0232 | -104.0449
25.027 | -104.0622
25.0307 | -104.0822
25.0344 | -104.1033
25.0381 | -104.1269
25.0418 | -104.1509
25.0456 | -104.1755
25.0493 | -104.2019
25.053 | -104.2269
25.0567 | -104.2521
25.0604 | -104.2792
25.0642 | -104.3058
25.0679 | -104.3329
25.0716 | -104.3616
25.0753 | -104.3908
25.079 | -104.4197
25.0828 | -104.448
25.0865 | -104.4768
25.0902 | -104.5043
25.0939 | -104.5305
25.0976 | -104.5551
25.1014 | -104.581
25.1051 | -104.6046
25.1088 | -104.6259
25.1125 | -104.6466
25.1162 | -104.6676
25.1199 | -104.6886
25.1237 | -104.7095
25.1274 | -104.7305
25.1311 | -104.748
25.1348 | -104.7643
25.1385 | -104.7816
25.1423 | -104.7972
25.146 | -104.8109
25.1497 | -104.8229
25.1534 | -104.8345
25.1571 | -104.8457
25.1609 | -104.8587
25.1646 | -104.8707
25.1683 | -104.8845
25.172 | -104.9017
25.1757 | -104.922
25.1795 | -104.9439
25.1832 | -104.9536
25.1869 | -104.9564
25.1906 | -104.9586
25.1943 | -104.9589
25.1981 | -104.9601
25.2018 | -104.9614
25.2055 | -104.9632
25.2092 | -104.9647
25.2129 | -104.9658
25.2166 | -104.9658
25.2204 | -104.966
25.2241 | -104.9661
25.2278 | -104.9667
25.2315 | -104.9667
25.2352 | -104.9661
25.239 | -104.9649
25.2427 | -104.965
25.2464 | -104.9651
25.2501 | -104.9643
25.2538 | -104.9639
25.2576 | -104.9648
25.2613 | -104.9647
25.265 | -104.9632
25.2687 | -104.9615
25.2724 | -104.9612
25.2762 | -104.9616
25.2799 | -104.9605
25.2836 | -104.9593
25.2873 | -104.9601
25.291 | -104.9605
25.2948 | -104.9606
25.2985 | -104.9587
25.3022 | -104.9572
25.3059 | -104.9588
25.3096 | -104.9616
25.3133 | -104.9637
25.3171 | -104.9649
25.3208 | -104.9677
25.3245 | -104.9717
25.3282 | -104.9735
25.3319 | -104.9735
25.3357 | -104.9734
25.3394 | -104.9743
25.3431 | -104.9755
25.3468 | -104.978
25.3505 | -104.9788
25.3543 | -104.979
25.358 | -104.9816
25.3617 | -104.984
25.3654 | -104.9875
25.3691 | -104.9886
25.3729 | -104.9874
25.3766 | -104.9864
25.3803 | -104.988
25.384 | -104.9901
25.3877 | -104.9916
25.3915 | -104.9929
25.3952 | -104.9932
25.3989 | -104.9927
25.4026 | -104.9921
25.4063 | -104.9928
25.41 | -104.9931
25.4138 | -104.9934
25.4175 | -104.993
25.4212 | -104.9926
25.4249 | -104.991
25.4286 | -104.9874
25.4324 | -104.9828
25.4361 | -104.9784
25.4398 | -104.9745
25.4435 | -104.9707
25.4472 | -104.9668
25.451 | -104.964
25.4547 | -104.9608
25.4584 | -104.9576
25.4621 | -104.9542
25.4658 | -104.9513
25.4696 | -104.9478
25.4733 | -104.9419
25.477 | -104.9352
25.4807 | -104.9284
25.4844 | -104.9223
25.4882 | -104.9164
25.4919 | -104.91
25.4956 | -104.9028
25.4993 | -104.8956
25.503 | -104.8875
25.5068 | -104.8789
25.5105 | -104.8697
25.5142 | -104.8612
25.5179 | -104.852
25.5216 | -104.8429
25.5253 | -104.8349
25.5291 | -104.8275
25.5328 | -104.8201
25.5365 | -104.8139
25.5402 | -104.8076
25.5439 | -104.8009
25.5477 | -104.7937
25.5514 | -104.7857
25.5551 | -104.7775
25.5588 | -104.7711
25.5625 | -104.766
25.5663 | -104.759
25.57 | -104.7529
25.5737 | -104.7464
25.5774 | -104.7404
25.5811 | -104.7353
25.5849 | -104.7305
25.5886 | -104.7251
25.5923 | -104.7185
25.596 | -104.7142
25.5997 | -104.7115
25.6035 | -104.7089
25.6072 | -104.7052
25.6109 | -104.7012
25.6146 | -104.6965
25.6183 | -104.6927
25.622 | -104.6906
25.6258 | -104.6884
25.6295 | -104.6856
25.6332 | -104.6791
25.6369 | -104.6732
25.6406 | -104.6684
25.6444 | -104.6636
25.6481 | -104.6595
25.6518 | -104.6556
25.6555 | -104.6525
25.6592 | -104.6497
25.663 | -104.6474
25.6667 | -104.644
25.6704 | -104.64
25.6741 | -104.6364
25.6778 | -104.6336
25.6816 | -104.6306
25.6853 | -104.6286
25.689 | -104.6254
25.6927 | -104.6239
25.6964 | -104.6229
25.7002 | -104.6224
25.7039 | -104.6217
25.7076 | -104.621
25.7113 | -104.6206
25.715 | -104.6207
25.7187 | -104.62
25.7225 | -104.6203
25.7262 | -104.6211
25.7299 | -104.625
25.7336 | -104.6304
25.7373 | -104.6362
25.7411 | -104.6411
25.7448 | -104.6446
25.7485 | -104.6491
25.7522 | -104.6614
25.7559 | -104.6669
25.7597 | -104.6732
25.7634 | -104.6809
25.7671 | -104.689
25.7708 | -104.6965
25.7745 | -104.7049
25.7783 | -104.7124
25.782 | -104.722
25.7857 | -104.732
25.7894 | -104.7424
25.7931 | -104.7539
25.7969 | -104.7648
25.8006 | -104.777
25.8043 | -104.7902
25.808 | -104.8029
25.8117 | -104.8157
25.8154 | -104.8279
25.8192 | -104.839
25.8229 | -104.8511
25.8266 | -104.864
25.8303 | -104.8771
25.834 | -104.8905
25.8378 | -104.9057
25.8415 | -104.921
25.8452 | -104.9361
25.8489 | -104.9508
25.8526 | -104.9634
25.8564 | -104.9763
25.8601 | -104.9903
25.8638 | -105.0052
25.8675 | -105.0204
25.8712 | -105.0337
25.875 | -105.0483
25.8787 | -105.0621
25.8824 | -105.0753
25.8861 | -105.0889
25.8898 | -105.0992
25.8936 | -105.1089
25.8973 | -105.118
25.901 | -105.126
25.9047 | -105.1363
25.9084 | -105.1446
25.9122 | -105.1521
25.9159 | -105.1613
25.9196 | -105.1685
25.9233 | -105.1748
25.927 | -105.1804
25.9307 | -105.1868
25.9345 | -105.1904
25.9382 | -105.1924
25.9419 | -105.194
25.9456 | -105.1955
25.9493 | -105.1969
25.9531 | -105.1985
25.9568 | -105.2008
25.9605 | -105.1998
25.9642 | -105.1961
25.9679 | -105.1931
25.9717 | -105.1927
25.9754 | -105.1919
25.9791 | -105.1905
25.9828 | -105.1873
25.9865 | -105.1837
25.9903 | -105.1828
25.994 | -105.1832
25.9977 | -105.1819
26.0014 | -105.1801
26.0051 | -105.1804
26.0089 | -105.1791
26.0126 | -105.1763
26.0163 | -105.1739
26.02 | -105.1726
26.0237 | -105.1726
26.0274 | -105.173
26.0312 | -105.1733
26.0349 | -105.1742
26.0386 | -105.1752
26.0423 | -105.176
26.046 | -105.1769
26.0498 | -105.1786
26.0535 | -105.1818
26.0572 | -105.1848
26.0609 | -105.1851
26.0646 | -105.1874
26.0684 | -105.1901
26.0721 | -105.1909
26.0758 | -105.193
26.0795 | -105.1945
26.0832 | -105.1972
26.087 | -105.2002
26.0907 | -105.2027
26.0944 | -105.2043
26.0981 | -105.2057
26.1018 | -105.2076
26.1056 | -105.2096
26.1093 | -105.2123
26.113 | -105.2149
26.1167 | -105.2153
26.1204 | -105.2169
26.1241 | -105.2202
26.1279 | -105.2225
26.1316 | -105.2249
26.1353 | -105.2279
26.139 | -105.2311
26.1427 | -105.235
26.1465 | -105.2395
26.1502 | -105.2417
26.1539 | -105.2427
26.1576 | -105.2432
26.1613 | -105.2454
26.1651 | -105.249
26.1688 | -105.2514
26.1725 | -105.2527
26.1762 | -105.2544
26.1799 | -105.2574
26.1837 | -105.261
26.1874 | -105.2635
26.1911 | -105.2653
26.1948 | -105.2659
26.1985 | -105.2689
26.2023 | -105.2721
26.206 | -105.274
26.2097 | -105.2756
26.2134 | -105.2765
26.2171 | -105.2765
26.2208 | -105.279
26.2246 | -105.2816
26.2283 | -105.2834
26.232 | -105.2858
26.2357 | -105.2883
26.2394 | -105.2929
26.2432 | -105.2955
26.2469 | -105.297
26.2506 | -105.2981
26.2543 | -105.2981
26.258 | -105.2994
26.2618 | -105.3008
26.2655 | -105.3031
26.2692 | -105.3049
26.2729 | -105.3055
26.2766 | -105.3065
26.2804 | -105.3073
26.2841 | -105.3066
26.2878 | -105.3074
26.2915 | -105.3069
26.2952 | -105.3062
26.299 | -105.3053
26.3027 | -105.3051
26.3064 | -105.3058
26.3101 | -105.3066
26.3138 | -105.3072
26.3176 | -105.307
26.3213 | -105.3064
26.325 | -105.3057
26.3287 | -105.3045
26.3324 | -105.3046
26.3361 | -105.3069
26.3399 | -105.3067
26.3436 | -105.3059
26.3473 | -105.3055
26.351 | -105.3058
26.3547 | -105.3071
26.3585 | -105.3073
26.3622 | -105.3071
26.3659 | -105.3064
26.3696 | -105.308
26.3733 | -105.3086
26.3771 | -105.3071
26.3808 | -105.3078
26.3845 | -105.3087
26.3882 | -105.3089
26.3919 | -105.3092
26.3957 | -105.3098
26.3994 | -105.3101
26.4031 | -105.3118
26.4068 | -105.3127
26.4105 | -105.313
26.4143 | -105.3157
26.418 | -105.3206
26.4217 | -105.3227
26.4254 | -105.3242
26.4291 | -105.3263
26.4328 | -105.3309
26.4366 | -105.3337
26.4403 | -105.3353
26.444 | -105.3394
26.4477 | -105.3433
26.4514 | -105.3473
26.4552 | -105.3504
26.4589 | -105.3518
26.4626 | -105.3553
26.4663 | -105.359
26.47 | -105.3635
26.4738 | -105.3675
26.4775 | -105.3726
26.4812 | -105.3761
26.4849 | -105.3789
26.4886 | -105.3826
26.4924 | -105.3854
26.4961 | -105.3876
26.4998 | -105.3908
26.5035 | -105.3942
26.5072 | -105.3975
26.511 | -105.3996
26.5147 | -105.4021
26.5184 | -105.4062
26.5221 | -105.4103
26.5258 | -105.4154
26.5296 | -105.4219
26.5333 | -105.4259
26.537 | -105.43
26.5407 | -105.4344
26.5444 | -105.4385
26.5481 | -105.4431
26.5519 | -105.4484
26.5556 | -105.4528
26.5593 | -105.4573
26.563 | -105.4618
26.5667 | -105.4686
26.5705 | -105.4724
26.5742 | -105.4766
26.5779 | -105.4827
26.5816 | -105.4884
26.5853 | -105.4946
26.5891 | -105.5009
26.5928 | -105.5076
26.5965 | -105.5145
26.6002 | -105.5194
26.6039 | -105.524
26.6077 | -105.5297
26.6114 | -105.5338
26.6151 | -105.5368
26.6188 | -105.5413
26.6225 | -105.5471
26.6263 | -105.5532
26.63 | -105.5574
26.6337 | -105.5617
26.6374 | -105.5649
26.6411 | -105.5667
26.6448 | -105.5673
26.6486 | -105.5686
26.6523 | -105.5719
26.656 | -105.5761
26.6597 | -105.579
26.6634 | -105.58
26.6672 | -105.582
26.6709 | -105.5834
26.6746 | -105.5834
26.6783 | -105.582
26.682 | -105.5807
26.6858 | -105.5806
26.6895 | -105.5798
26.6932 | -105.5772
26.6969 | -105.574
26.7006 | -105.5718
26.7044 | -105.5699
26.7081 | -105.5658
26.7118 | -105.5604
26.7155 | -105.5568
26.7192 | -105.5518
26.723 | -105.5464
26.7267 | -105.5407
26.7304 | -105.534
26.7341 | -105.5284
26.7378 | -105.5219
26.7415 | -105.5154
26.7453 | -105.5092
26.749 | -105.5037
26.7527 | -105.4982
26.7564 | -105.4926
26.7601 | -105.4873
26.7639 | -105.4815
26.7676 | -105.4756
26.7713 | -105.4698
26.775 | -105.4637
26.7787 | -105.4581
26.7825 | -105.4533
26.7862 | -105.45
26.7899 | -105.4448
26.7936 | -105.4371
26.7973 | -105.4328
26.8011 | -105.4306
26.8048 | -105.4285
26.8085 | -105.4241
26.8122 | -105.4183
26.8159 | -105.4143
26.8197 | -105.4108
26.8234 | -105.4094
26.8271 | -105.4081
26.8308 | -105.406
26.8345 | -105.403
26.8382 | -105.4
26.842 | -105.3974
26.8457 | -105.3961
26.8494 | -105.3927
26.8531 | -105.3877
26.8568 | -105.3833
26.8606 | -105.3808
26.8643 | -105.3794
26.868 | -105.3782
26.8717 | -105.3765
26.8754 | -105.3734
26.8792 | -105.3704
26.8829 | -105.3664
26.8866 | -105.3606
26.8903 | -105.3568
26.894 | -105.3527
26.8978 | -105.3465
26.9015 | -105.3411
26.9052 | -105.3356
26.9089 | -105.3309
26.9126 | -105.3266
26.9164 | -105.3205
26.9201 | -105.3142
26.9238 | -105.3078
26.9275 | -105.3016
26.9312 | -105.2937
26.935 | -105.2862
26.9387 | -105.281
26.9424 | -105.2746
26.9461 | -105.267
26.9498 | -105.2605
26.9535 | -105.2546
26.9573 | -105.2475
26.961 | -105.2413
26.9647 | -105.2361
26.9684 | -105.2308
26.9721 | -105.2246
26.9759 | -105.2174
26.9796 | -105.2121
26.9833 | -105.2076
26.987 | -105.2028
26.9907 | -105.1988
26.9945 | -105.1955
26.9982 | -105.1912
27.0019 | -105.1873
27.0056 | -105.1839
27.0093 | -105.1812
27.0131 | -105.1788
27.0168 | -105.1776
27.0205 | -105.1779
27.0242 | -105.1771
27.0279 | -105.1766
27.0317 | -105.1749
27.0354 | -105.1733
27.0391 | -105.1737
27.0428 | -105.1744
27.0465 | -105.1755
27.0502 | -105.1774
27.054 | -105.1797
27.0577 | -105.1833
27.0614 | -105.1852
27.0651 | -105.1875
27.0688 | -105.1908
27.0726 | -105.1946
27.0763 | -105.1979
27.08 | -105.2016
27.0837 | -105.2055
27.0874 | -105.21
27.0912 | -105.2148
27.0949 | -105.2192
27.0986 | -105.2238
27.1023 | -105.2286
27.106 | -105.2322
27.1098 | -105.2348
27.1135 | -105.2377
27.1172 | -105.2429
27.1209 | -105.2476
27.1246 | -105.2513
27.1284 | -105.2542
27.1321 | -105.2562
27.1358 | -105.2572
27.1395 | -105.2576
27.1432 | -105.2593
27.1469 | -105.2593
27.1507 | -105.2572
27.1544 | -105.2553
27.1581 | -105.2536
27.1618 | -105.2525
27.1655 | -105.2489
27.1693 | -105.245
27.173 | -105.2407
27.1767 | -105.2356
27.1804 | -105.2299
27.1841 | -105.2229
27.1879 | -105.2177
27.1916 | -105.2142
27.1953 | -105.2087
27.199 | -105.2008
27.2027 | -105.1909
27.2065 | -105.1824
27.2102 | -105.1748
27.2139 | -105.1649
27.2176 | -105.1563
27.2213 | -105.1476
27.2251 | -105.1372
27.2288 | -105.1272
27.2325 | -105.1187
27.2362 | -105.1106
27.2399 | -105.1009
27.2436 | -105.0915
27.2474 | -105.0805
27.2511 | -105.0695
27.2548 | -105.0588
27.2585 | -105.0507
27.2622 | -105.0415
27.266 | -105.0325
27.2697 | -105.0251
27.2734 | -105.0165
27.2771 | -105.0078
27.2808 | -105.0011
27.2846 | -104.9954
27.2883 | -104.9893
27.292 | -104.9823
27.2957 | -104.9765
27.2994 | -104.9728
27.3032 | -104.9692
27.3069 | -104.9663
27.3106 | -104.9625
27.3143 | -104.9603
27.318 | -104.9599
27.3218 | -104.9589
27.3255 | -104.9592
27.3292 | -104.9586
27.3329 | -104.9578
27.3366 | -104.9581
27.3404 | -104.96
27.3441 | -104.9626
27.3478 | -104.9656
27.3515 | -104.9689
27.3552 | -104.9739
27.3589 | -104.9785
27.3627 | -104.9824
27.3664 | -104.9858
27.3701 | -104.989
27.3738 | -104.9928
27.3775 | -104.9964
27.3813 | -105.0005
27.385 | -105.0059
27.3887 | -105.0093
27.3924 | -105.0129
27.3961 | -105.018
27.3999 | -105.0217
27.4036 | -105.0258
27.4073 | -105.0298
27.411 | -105.0344
27.4147 | -105.0395
27.4185 | -105.0432
27.4222 | -105.0466
27.4259 | -105.0485
27.4296 | -105.0513
27.4333 | -105.0532
27.4371 | -105.0548
27.4408 | -105.0562
27.4445 | -105.0578
27.4482 | -105.0606
27.4519 | -105.0631
27.4556 | -105.0637
27.4594 | -105.0647
27.4631 | -105.0648
27.4668 | -105.0634
27.4705 | -105.0616
27.4742 | -105.0613
27.478 | -105.0599
27.4817 | -105.0564
27.4854 | -105.0537
27.4891 | -105.0526
27.4928 | -105.0532
27.4966 | -105.0541
27.5003 | -105.0539
27.504 | -105.0526
27.5077 | -105.0509
27.5114 | -105.0488
27.5152 | -105.0451
27.5189 | -105.0429
27.5226 | -105.0423
27.5263 | -105.0424
27.53 | -105.0424
27.5338 | -105.0421
27.5375 | -105.0421
27.5412 | -105.0429
27.5449 | -105.0421
27.5486 | -105.0401
27.5523 | -105.0401
27.5561 | -105.042
27.5598 | -105.0442
27.5635 | -105.0454
27.5672 | -105.0461
27.5709 | -105.0468
27.5747 | -105.0474
27.5784 | -105.0469
27.5821 | -105.0461
27.5858 | -105.046
27.5895 | -105.0459
27.5933 | -105.0456
27.597 | -105.0464
27.6007 | -105.0467
27.6044 | -105.0469
27.6081 | -105.0489
27.6119 | -105.0515
27.6156 | -105.0529
27.6193 | -105.0534
27.623 | -105.0534
27.6267 | -105.0552
27.6305 | -105.0574
27.6342 | -105.0595
27.6379 | -105.0614
27.6416 | -105.0643
27.6453 | -105.0648
27.649 | -105.0656
27.6528 | -105.0679
27.6565 | -105.0695
27.6602 | -105.0715
27.6639 | -105.0744
27.6676 | -105.0751
27.6714 | -105.0749
27.6751 | -105.0759
27.6788 | -105.0778
27.6825 | -105.0784
27.6862 | -105.0801
27.69 | -105.0831
27.6937 | -105.0854
27.6974 | -105.0876
27.7011 | -105.0902
27.7048 | -105.0926
27.7086 | -105.0937
27.7123 | -105.094
27.716 | -105.0967
27.7197 | -105.0988
27.7234 | -105.0991
27.7272 | -105.099
27.7309 | -105.0985
27.7346 | -105.0983
27.7383 | -105.0975
27.742 | -105.0949
27.7458 | -105.0941
27.7495 | -105.0945
27.7532 | -105.0952
27.7569 | -105.095
27.7606 | -105.0942
27.7643 | -105.0939
27.7681 | -105.0939
27.7718 | -105.0924
27.7755 | -105.0901
27.7792 | -105.088
27.7829 | -105.0854
27.7867 | -105.0826
27.7904 | -105.0803
27.7941 | -105.0755
27.7978 | -105.0704
27.8015 | -105.0675
27.8053 | -105.0633
27.809 | -105.0608
27.8127 | -105.0581
27.8164 | -105.0555
27.8201 | -105.051
27.8239 | -105.0464
27.8276 | -105.0416
27.8313 | -105.0377
27.835 | -105.0332
27.8387 | -105.0294
27.8425 | -105.0247
27.8462 | -105.0211
27.8499 | -105.0183
27.8536 | -105.0172
27.8573 | -105.0173
27.861 | -105.0179
27.8648 | -105.0167
27.8685 | -105.0166
27.8722 | -105.018
27.8759 | -105.0183
27.8796 | -105.0177
27.8834 | -105.0155
27.8871 | -105.0137
27.8908 | -105.012
27.8945 | -105.0107
27.8982 | -105.013
27.902 | -105.0127
27.9057 | -105.0137
27.9094 | -105.0155
27.9131 | -105.0165
27.9168 | -105.017
27.9206 | -105.0174
27.9243 | -105.0174
27.928 | -105.0178
27.9317 | -105.0186
27.9354 | -105.0189
27.9392 | -105.0197
27.9429 | -105.0212
27.9466 | -105.0227
27.9503 | -105.0223
27.954 | -105.0229
27.9577 | -105.023
27.9615 | -105.0231
27.9652 | -105.0235
27.9689 | -105.0225
27.9726 | -105.0225
27.9763 | -105.024
27.9801 | -105.0253
27.9838 | -105.0252
27.9875 | -105.0247
27.9912 | -105.0225
27.9949 | -105.0214
27.9987 | -105.0206
28.0024 | -105.0191