## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3d | 28 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 16
Hood | 02-91C/P | 17
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1001.1 | hPa |  | PRC PRO 01
Humidity | 31.5 | %/Hr
Internal temperature of calibration bench | 23.43 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1003 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 23.5 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 1.99 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.77 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.92 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3187 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Analog acquisition system | Analog Hood
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 09/03/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 13/03/2018
Technical Manager | Anthony HUE | Approval |  | 20/03/2018
CEO | Robert GLEMOT | Approval |  | 20/03/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.399 | 57.3692
0.012209 | -35.9524 | 38.2162
0.018313 | -34.9033 | 28.3369
0.024417 | -34.4619 | 21.9334
0.030521 | -34.3557 | 18.1876
0.036626 | -34.2184 | 15.1446
0.04273 | -34.1348 | 12.8415
0.048834 | -34.0845 | 11.2693
0.054938 | -34.0186 | 10.0585
0.061043 | -33.9474 | 8.9062
0.067147 | -33.9358 | 8.0222
0.073251 | -33.9356 | 7.2588
0.079355 | -33.9527 | 6.6882
0.08546 | -33.97 | 6.2165
0.091564 | -33.9663 | 5.8387
0.097668 | -33.9346 | 5.5852
0.10377 | -33.8935 | 5.2465
0.10988 | -33.8661 | 4.8937
0.11598 | -33.8567 | 4.5463
0.12209 | -33.8696 | 4.2753
0.12819 | -33.8979 | 4.0445
0.13429 | -33.9279 | 3.8005
0.1404 | -33.9464 | 3.5186
0.1465 | -33.9627 | 3.2554
0.15261 | -33.9709 | 3.0135
0.15871 | -33.9686 | 2.8188
0.16482 | -33.9554 | 2.6634
0.17092 | -33.9329 | 2.5949
0.17702 | -33.9157 | 2.4932
0.18313 | -33.9002 | 2.403
0.18923 | -33.8893 | 2.2906
0.19534 | -33.8832 | 2.1811
0.20144 | -33.8823 | 2.0627
0.20754 | -33.8839 | 1.9853
0.21365 | -33.8854 | 1.9058
0.21975 | -33.8886 | 1.8379
0.22586 | -33.8954 | 1.7704
0.23196 | -33.903 | 1.6976
0.23807 | -33.9111 | 1.6208
0.24417 | -33.914 | 1.5327
0.25027 | -33.9177 | 1.4519
0.25638 | -33.9184 | 1.3652
0.26248 | -33.9174 | 1.2816
0.26859 | -33.9159 | 1.201
0.27469 | -33.9157 | 1.1167
0.2808 | -33.921 | 1.0499
0.2869 | -33.9252 | 0.97967
0.293 | -33.9303 | 0.91915
0.29911 | -33.9359 | 0.87539
0.30521 | -33.9414 | 0.83225
0.31132 | -33.9451 | 0.81158
0.31742 | -33.9489 | 0.78343
0.32353 | -33.952 | 0.76016
0.32963 | -33.9547 | 0.74048
0.33573 | -33.9563 | 0.72423
0.34184 | -33.9587 | 0.70596
0.34794 | -33.9617 | 0.67174
0.35405 | -33.9649 | 0.6418
0.36015 | -33.9683 | 0.60875
0.36626 | -33.9707 | 0.57187
0.37236 | -33.9732 | 0.52331
0.37846 | -33.9755 | 0.46744
0.38457 | -33.9747 | 0.40707
0.39067 | -33.9754 | 0.35621
0.39678 | -33.9739 | 0.29852
0.40288 | -33.9717 | 0.24287
0.40899 | -33.9682 | 0.18833
0.41509 | -33.9639 | 0.13299
0.42119 | -33.9593 | 0.080833
0.4273 | -33.9558 | 0.028004
0.4334 | -33.9527 | -0.025893
0.43951 | -33.95 | -0.07788
0.44561 | -33.9487 | -0.12701
0.45172 | -33.9502 | -0.16736
0.45782 | -33.95 | -0.21181
0.46392 | -33.9513 | -0.25208
0.47003 | -33.9528 | -0.28923
0.47613 | -33.9546 | -0.32296
0.48224 | -33.9565 | -0.35466
0.48834 | -33.9585 | -0.37589
0.49445 | -33.9602 | -0.40135
0.50055 | -33.9622 | -0.42439
0.50665 | -33.9637 | -0.4431
0.51276 | -33.9645 | -0.46289
0.51886 | -33.9655 | -0.48083
0.52497 | -33.9646 | -0.48876
0.53107 | -33.9649 | -0.50047
0.53717 | -33.9652 | -0.51255
0.54328 | -33.965 | -0.5194
0.54938 | -33.9651 | -0.52935
0.55549 | -33.9647 | -0.53963
0.56159 | -33.9648 | -0.55234
0.5677 | -33.9644 | -0.5646
0.5738 | -33.9641 | -0.57959
0.5799 | -33.9642 | -0.59325
0.58601 | -33.9647 | -0.61092
0.59211 | -33.9657 | -0.63506
0.59822 | -33.9663 | -0.65393
0.60432 | -33.9674 | -0.67717
0.61043 | -33.9683 | -0.69988
0.61653 | -33.9694 | -0.72358
0.62263 | -33.9704 | -0.74881
0.62874 | -33.9716 | -0.77222
0.63484 | -33.9726 | -0.79625
0.64095 | -33.9737 | -0.82058
0.64705 | -33.9748 | -0.84338
0.65316 | -33.9756 | -0.86613
0.65926 | -33.9765 | -0.8922
0.66536 | -33.9761 | -0.91408
0.67147 | -33.9764 | -0.93405
0.67757 | -33.9769 | -0.9554
0.68368 | -33.9768 | -0.97815
0.68978 | -33.9771 | -0.99925
0.69589 | -33.9769 | -1.02
0.70199 | -33.9765 | -1.0439
0.70809 | -33.9761 | -1.0666
0.7142 | -33.9759 | -1.0881
0.7203 | -33.9754 | -1.1097
0.72641 | -33.975 | -1.1324
0.73251 | -33.9751 | -1.1548
0.73862 | -33.9755 | -1.1794
0.74472 | -33.9761 | -1.2003
0.75082 | -33.976 | -1.2228
0.75693 | -33.9765 | -1.2472
0.76303 | -33.9768 | -1.2706
0.76914 | -33.9775 | -1.2979
0.77524 | -33.978 | -1.3235
0.78135 | -33.9787 | -1.3479
0.78745 | -33.9795 | -1.3737
0.79355 | -33.9802 | -1.3988
0.79966 | -33.9812 | -1.4254
0.80576 | -33.9824 | -1.4563
0.81187 | -33.9829 | -1.4864
0.81797 | -33.9836 | -1.5171
0.82408 | -33.9843 | -1.5448
0.83018 | -33.985 | -1.5764
0.83628 | -33.9858 | -1.6072
0.84239 | -33.9857 | -1.6364
0.84849 | -33.9858 | -1.6653
0.8546 | -33.9863 | -1.6935
0.8607 | -33.9864 | -1.7233
0.86681 | -33.9865 | -1.7524
0.87291 | -33.9861 | -1.7792
0.87901 | -33.9857 | -1.8031
0.88512 | -33.9856 | -1.8275
0.89122 | -33.9855 | -1.8518
0.89733 | -33.9853 | -1.8726
0.90343 | -33.985 | -1.8955
0.90953 | -33.9852 | -1.9118
0.91564 | -33.9846 | -1.9294
0.92174 | -33.9843 | -1.9449
0.92785 | -33.9838 | -1.9601
0.93395 | -33.9837 | -1.9738
0.94006 | -33.9837 | -1.9887
0.94616 | -33.9835 | -1.9951
0.95226 | -33.9835 | -2.008
0.95837 | -33.9836 | -2.021
0.96447 | -33.9839 | -2.0328
0.97058 | -33.9838 | -2.0443
0.97668 | -33.9835 | -2.054
0.98279 | -33.9835 | -2.0598
0.98889 | -33.9836 | -2.0666
0.99499 | -33.9836 | -2.0767
1.0011 | -33.9838 | -2.0836
1.0072 | -33.9836 | -2.0892
1.0133 | -33.9835 | -2.0951
1.0194 | -33.9834 | -2.1028
1.0255 | -33.9831 | -2.1115
1.0316 | -33.9825 | -2.118
1.0377 | -33.9819 | -2.1262
1.0438 | -33.9812 | -2.1329
1.0499 | -33.9798 | -2.1417
1.056 | -33.9786 | -2.1493
1.0621 | -33.9776 | -2.1576
1.0682 | -33.9771 | -2.1654
1.0743 | -33.976 | -2.1754
1.0805 | -33.9754 | -2.1849
1.0866 | -33.9738 | -2.1989
1.0927 | -33.9725 | -2.209
1.0988 | -33.9717 | -2.2224
1.1049 | -33.9703 | -2.2326
1.111 | -33.9688 | -2.2467
1.1171 | -33.9675 | -2.2591
1.1232 | -33.9662 | -2.276
1.1293 | -33.9651 | -2.2903
1.1354 | -33.9642 | -2.305
1.1415 | -33.963 | -2.3214
1.1476 | -33.9617 | -2.3376
1.1537 | -33.9604 | -2.3531
1.1598 | -33.9603 | -2.3734
1.1659 | -33.9598 | -2.388
1.172 | -33.9591 | -2.4034
1.1781 | -33.9583 | -2.4199
1.1842 | -33.9577 | -2.4379
1.1903 | -33.9584 | -2.4552
1.1964 | -33.9582 | -2.4717
1.2025 | -33.9582 | -2.4891
1.2086 | -33.9586 | -2.5064
1.2147 | -33.9588 | -2.5231
1.2209 | -33.9596 | -2.5385
1.227 | -33.9614 | -2.555
1.2331 | -33.9623 | -2.5711
1.2392 | -33.9631 | -2.5884
1.2453 | -33.9643 | -2.6021
1.2514 | -33.9656 | -2.6158
1.2575 | -33.9671 | -2.6311
1.2636 | -33.9693 | -2.6438
1.2697 | -33.9709 | -2.657
1.2758 | -33.9725 | -2.6724
1.2819 | -33.9742 | -2.6875
1.288 | -33.9758 | -2.7027
1.2941 | -33.9778 | -2.7147
1.3002 | -33.9797 | -2.7247
1.3063 | -33.9813 | -2.7363
1.3124 | -33.9831 | -2.7479
1.3185 | -33.985 | -2.7604
1.3246 | -33.9868 | -2.7729
1.3307 | -33.9889 | -2.7843
1.3368 | -33.9908 | -2.7937
1.3429 | -33.9925 | -2.8047
1.349 | -33.994 | -2.8151
1.3551 | -33.9956 | -2.8281
1.3613 | -33.9971 | -2.838
1.3674 | -33.9984 | -2.8485
1.3735 | -33.9997 | -2.8586
1.3796 | -34.0008 | -2.8682
1.3857 | -34.0019 | -2.8777
1.3918 | -34.003 | -2.8895
1.3979 | -34.0043 | -2.9005
1.404 | -34.005 | -2.9133
1.4101 | -34.006 | -2.9248
1.4162 | -34.0069 | -2.9356
1.4223 | -34.0077 | -2.9485
1.4284 | -34.0085 | -2.9594
1.4345 | -34.0091 | -2.9724
1.4406 | -34.0096 | -2.9859
1.4467 | -34.0104 | -2.9981
1.4528 | -34.0111 | -3.0096
1.4589 | -34.012 | -3.0224
1.465 | -34.0123 | -3.0331
1.4711 | -34.0132 | -3.0451
1.4772 | -34.0138 | -3.0601
1.4833 | -34.0146 | -3.0748
1.4894 | -34.0153 | -3.0887
1.4955 | -34.016 | -3.103
1.5016 | -34.0168 | -3.1179
1.5078 | -34.0173 | -3.1352
1.5139 | -34.0181 | -3.1504
1.52 | -34.0187 | -3.1659
1.5261 | -34.0191 | -3.1802
1.5322 | -34.0194 | -3.1954
1.5383 | -34.0199 | -3.2105
1.5444 | -34.0207 | -3.2265
1.5505 | -34.0212 | -3.2422
1.5566 | -34.0218 | -3.2576
1.5627 | -34.0225 | -3.2734
1.5688 | -34.0234 | -3.2893
1.5749 | -34.0238 | -3.3043
1.581 | -34.0243 | -3.3231
1.5871 | -34.025 | -3.3414
1.5932 | -34.0256 | -3.3568
1.5993 | -34.0261 | -3.3758
1.6054 | -34.0266 | -3.3934
1.6115 | -34.0271 | -3.4128
1.6176 | -34.0277 | -3.4314
1.6237 | -34.028 | -3.4482
1.6298 | -34.0285 | -3.466
1.6359 | -34.0291 | -3.4831
1.642 | -34.0297 | -3.4999
1.6482 | -34.0301 | -3.5185
1.6543 | -34.0308 | -3.5377
1.6604 | -34.0313 | -3.554
1.6665 | -34.0316 | -3.5714
1.6726 | -34.0321 | -3.5896
1.6787 | -34.0327 | -3.6089
1.6848 | -34.0327 | -3.6273
1.6909 | -34.0332 | -3.6451
1.697 | -34.0336 | -3.6632
1.7031 | -34.0338 | -3.6804
1.7092 | -34.0339 | -3.6971
1.7153 | -34.0342 | -3.7147
1.7214 | -34.0339 | -3.7311
1.7275 | -34.0338 | -3.7489
1.7336 | -34.0337 | -3.7671
1.7397 | -34.0339 | -3.7838
1.7458 | -34.034 | -3.802
1.7519 | -34.0339 | -3.8198
1.758 | -34.0339 | -3.8357
1.7641 | -34.034 | -3.8518
1.7702 | -34.0339 | -3.869
1.7763 | -34.0339 | -3.8842
1.7824 | -34.0338 | -3.9018
1.7885 | -34.0337 | -3.9171
1.7947 | -34.0338 | -3.9334
1.8008 | -34.0338 | -3.9477
1.8069 | -34.0337 | -3.9629
1.813 | -34.0338 | -3.9786
1.8191 | -34.0335 | -3.9942
1.8252 | -34.0334 | -4.0094
1.8313 | -34.0334 | -4.0251
1.8374 | -34.0334 | -4.04
1.8435 | -34.0335 | -4.0537
1.8496 | -34.0333 | -4.0695
1.8557 | -34.0332 | -4.0855
1.8618 | -34.0332 | -4.0987
1.8679 | -34.0332 | -4.1132
1.874 | -34.0333 | -4.1279
1.8801 | -34.0332 | -4.1424
1.8862 | -34.0334 | -4.1564
1.8923 | -34.0333 | -4.1702
1.8984 | -34.0331 | -4.1823
1.9045 | -34.0331 | -4.1947
1.9106 | -34.0328 | -4.2082
1.9167 | -34.0326 | -4.2213
1.9228 | -34.0326 | -4.2338
1.9289 | -34.0323 | -4.2443
1.9351 | -34.0324 | -4.2551
1.9412 | -34.0325 | -4.2667
1.9473 | -34.0326 | -4.2802
1.9534 | -34.0326 | -4.2933
1.9595 | -34.0325 | -4.3066
1.9656 | -34.0322 | -4.3188
1.9717 | -34.032 | -4.3309
1.9778 | -34.0322 | -4.3441
1.9839 | -34.0322 | -4.3568
1.99 | -34.0323 | -4.3687
1.9961 | -34.0321 | -4.3804
2.0022 | -34.0322 | -4.392
2.0083 | -34.0319 | -4.4037
2.0144 | -34.0319 | -4.4152
2.0205 | -34.0319 | -4.4284
2.0266 | -34.0321 | -4.442
2.0327 | -34.0322 | -4.4547
2.0388 | -34.0325 | -4.4697
2.0449 | -34.0326 | -4.4819
2.051 | -34.0328 | -4.4934
2.0571 | -34.0327 | -4.5066
2.0632 | -34.0327 | -4.5192
2.0693 | -34.0328 | -4.5318
2.0754 | -34.0331 | -4.5454
2.0816 | -34.0333 | -4.5592
2.0877 | -34.0334 | -4.5721
2.0938 | -34.0337 | -4.586
2.0999 | -34.0337 | -4.5991
2.106 | -34.034 | -4.6128
2.1121 | -34.0341 | -4.6268
2.1182 | -34.0343 | -4.6386
2.1243 | -34.0346 | -4.6509
2.1304 | -34.0348 | -4.663
2.1365 | -34.0352 | -4.6762
2.1426 | -34.0354 | -4.691
2.1487 | -34.0359 | -4.7044
2.1548 | -34.0364 | -4.719
2.1609 | -34.0371 | -4.7327
2.167 | -34.0376 | -4.7469
2.1731 | -34.038 | -4.7607
2.1792 | -34.0387 | -4.7736
2.1853 | -34.0393 | -4.7881
2.1914 | -34.0399 | -4.802
2.1975 | -34.0405 | -4.8159
2.2036 | -34.041 | -4.8306
2.2097 | -34.0414 | -4.8442
2.2158 | -34.0424 | -4.8561
2.222 | -34.043 | -4.8683
2.2281 | -34.0434 | -4.8807
2.2342 | -34.0439 | -4.8955
2.2403 | -34.0444 | -4.9075
2.2464 | -34.0452 | -4.9204
2.2525 | -34.0456 | -4.9329
2.2586 | -34.0465 | -4.9445
2.2647 | -34.0473 | -4.9558
2.2708 | -34.048 | -4.969
2.2769 | -34.0488 | -4.9823
2.283 | -34.0493 | -4.9939
2.2891 | -34.0499 | -5.006
2.2952 | -34.0505 | -5.0176
2.3013 | -34.0512 | -5.0298
2.3074 | -34.0519 | -5.0413
2.3135 | -34.0524 | -5.0545
2.3196 | -34.053 | -5.0655
2.3257 | -34.0533 | -5.0781
2.3318 | -34.0536 | -5.0893
2.3379 | -34.0542 | -5.1024
2.344 | -34.0546 | -5.1151
2.3501 | -34.055 | -5.128
2.3562 | -34.0554 | -5.1385
2.3623 | -34.0557 | -5.1495
2.3685 | -34.056 | -5.1604
2.3746 | -34.0563 | -5.1731
2.3807 | -34.0567 | -5.1839
2.3868 | -34.0569 | -5.1952
2.3929 | -34.0573 | -5.2067
2.399 | -34.0576 | -5.2167
2.4051 | -34.0581 | -5.2282
2.4112 | -34.0585 | -5.2394
2.4173 | -34.0588 | -5.2507
2.4234 | -34.0588 | -5.2615
2.4295 | -34.0591 | -5.2722
2.4356 | -34.0593 | -5.284
2.4417 | -34.0594 | -5.2961
2.4478 | -34.0595 | -5.3081
2.4539 | -34.0598 | -5.3184
2.46 | -34.0596 | -5.3305
2.4661 | -34.0596 | -5.3422
2.4722 | -34.0599 | -5.353
2.4783 | -34.06 | -5.3636
2.4844 | -34.0601 | -5.3747
2.4905 | -34.06 | -5.3861
2.4966 | -34.0598 | -5.3974
2.5027 | -34.0598 | -5.4077
2.5089 | -34.0598 | -5.4187
2.515 | -34.0597 | -5.4304
2.5211 | -34.0597 | -5.4411
2.5272 | -34.0594 | -5.4526
2.5333 | -34.0593 | -5.4646
2.5394 | -34.0594 | -5.4761
2.5455 | -34.0593 | -5.4877
2.5516 | -34.0592 | -5.499
2.5577 | -34.0592 | -5.5114
2.5638 | -34.0591 | -5.5241
2.5699 | -34.059 | -5.5366
2.576 | -34.0588 | -5.5492
2.5821 | -34.0587 | -5.5614
2.5882 | -34.0586 | -5.5731
2.5943 | -34.0585 | -5.5854
2.6004 | -34.0584 | -5.598
2.6065 | -34.0583 | -5.6086
2.6126 | -34.0583 | -5.6202
2.6187 | -34.0582 | -5.6314
2.6248 | -34.0582 | -5.643
2.6309 | -34.0581 | -5.6566
2.637 | -34.0583 | -5.67
2.6431 | -34.0583 | -5.682
2.6492 | -34.0582 | -5.6944
2.6554 | -34.0583 | -5.707
2.6615 | -34.0582 | -5.7192
2.6676 | -34.0582 | -5.7318
2.6737 | -34.0584 | -5.7451
2.6798 | -34.0585 | -5.7574
2.6859 | -34.0586 | -5.7701
2.692 | -34.0586 | -5.7823
2.6981 | -34.0586 | -5.7955
2.7042 | -34.0588 | -5.8098
2.7103 | -34.0588 | -5.8223
2.7164 | -34.0591 | -5.8355
2.7225 | -34.0593 | -5.8492
2.7286 | -34.0594 | -5.8626
2.7347 | -34.0595 | -5.8749
2.7408 | -34.0599 | -5.8867
2.7469 | -34.0599 | -5.9002
2.753 | -34.0603 | -5.9133
2.7591 | -34.0605 | -5.9276
2.7652 | -34.0606 | -5.9415
2.7713 | -34.0607 | -5.9539
2.7774 | -34.0612 | -5.9676
2.7835 | -34.0614 | -5.9806
2.7896 | -34.0616 | -5.9922
2.7958 | -34.062 | -6.0048
2.8019 | -34.0625 | -6.0171
2.808 | -34.0629 | -6.0313
2.8141 | -34.0633 | -6.0465
2.8202 | -34.0638 | -6.0609
2.8263 | -34.0642 | -6.0742
2.8324 | -34.0646 | -6.0889
2.8385 | -34.0651 | -6.103
2.8446 | -34.0657 | -6.117
2.8507 | -34.0661 | -6.1315
2.8568 | -34.0666 | -6.145
2.8629 | -34.067 | -6.1585
2.869 | -34.0675 | -6.1721
2.8751 | -34.068 | -6.1864
2.8812 | -34.0687 | -6.2019
2.8873 | -34.0691 | -6.2156
2.8934 | -34.0695 | -6.2295
2.8995 | -34.0698 | -6.2437
2.9056 | -34.0702 | -6.2569
2.9117 | -34.0708 | -6.2709
2.9178 | -34.0714 | -6.2856
2.9239 | -34.0719 | -6.301
2.93 | -34.0724 | -6.3168
2.9361 | -34.0729 | -6.3306
2.9423 | -34.0735 | -6.3448
2.9484 | -34.074 | -6.3604
2.9545 | -34.0748 | -6.3755
2.9606 | -34.0753 | -6.3908
2.9667 | -34.0759 | -6.4047
2.9728 | -34.0763 | -6.4191
2.9789 | -34.0768 | -6.434
2.985 | -34.0773 | -6.4497
2.9911 | -34.0778 | -6.4643
2.9972 | -34.0784 | -6.4792
3.0033 | -34.079 | -6.4944
3.0094 | -34.0796 | -6.5094
3.0155 | -34.0802 | -6.5241
3.0216 | -34.0807 | -6.5394
3.0277 | -34.0812 | -6.5552
3.0338 | -34.0818 | -6.5707
3.0399 | -34.0825 | -6.5864
3.046 | -34.0828 | -6.6015
3.0521 | -34.0834 | -6.616
3.0582 | -34.0838 | -6.6317
3.0643 | -34.0843 | -6.6468
3.0704 | -34.0846 | -6.6621
3.0765 | -34.0851 | -6.6764
3.0827 | -34.0854 | -6.6914
3.0888 | -34.0858 | -6.7063
3.0949 | -34.0862 | -6.7211
3.101 | -34.0867 | -6.737
3.1071 | -34.087 | -6.753
3.1132 | -34.0876 | -6.7689
3.1193 | -34.088 | -6.7838
3.1254 | -34.0885 | -6.7979
3.1315 | -34.0889 | -6.814
3.1376 | -34.0892 | -6.8287
3.1437 | -34.0897 | -6.8442
3.1498 | -34.0901 | -6.8602
3.1559 | -34.0905 | -6.8754
3.162 | -34.0906 | -6.8913
3.1681 | -34.0908 | -6.9073
3.1742 | -34.0912 | -6.9228
3.1803 | -34.0916 | -6.9379
3.1864 | -34.0921 | -6.9519
3.1925 | -34.0924 | -6.9674
3.1986 | -34.0926 | -6.9818
3.2047 | -34.0928 | -6.9975
3.2108 | -34.0931 | -7.0113
3.2169 | -34.0936 | -7.0269
3.223 | -34.0937 | -7.0421
3.2292 | -34.0938 | -7.0577
3.2353 | -34.0937 | -7.0691
3.2414 | -34.094 | -7.0839
3.2475 | -34.0941 | -7.0967
3.2536 | -34.0944 | -7.1122
3.2597 | -34.0947 | -7.1252
3.2658 | -34.0949 | -7.138
3.2719 | -34.0951 | -7.1511
3.278 | -34.0952 | -7.1652
3.2841 | -34.0955 | -7.1787
3.2902 | -34.0956 | -7.193
3.2963 | -34.0959 | -7.2063
3.3024 | -34.0958 | -7.2183
3.3085 | -34.0961 | -7.2325
3.3146 | -34.0962 | -7.2452
3.3207 | -34.0964 | -7.2592
3.3268 | -34.0966 | -7.2732
3.3329 | -34.0968 | -7.285
3.339 | -34.097 | -7.2965
3.3451 | -34.0971 | -7.3098
3.3512 | -34.0973 | -7.3211
3.3573 | -34.0976 | -7.3343
3.3634 | -34.0977 | -7.3475
3.3696 | -34.0978 | -7.3611
3.3757 | -34.0979 | -7.3727
3.3818 | -34.0982 | -7.3862
3.3879 | -34.0982 | -7.3985
3.394 | -34.0983 | -7.4109
3.4001 | -34.0984 | -7.4235
3.4062 | -34.0985 | -7.4356
3.4123 | -34.0987 | -7.4471
3.4184 | -34.0988 | -7.4599
3.4245 | -34.099 | -7.4723
3.4306 | -34.099 | -7.4839
3.4367 | -34.0992 | -7.4955
3.4428 | -34.0996 | -7.5065
3.4489 | -34.0997 | -7.5193
3.455 | -34.0997 | -7.5316
3.4611 | -34.0998 | -7.5445
3.4672 | -34.0999 | -7.5575
3.4733 | -34.1 | -7.5694
3.4794 | -34.1001 | -7.5817
3.4855 | -34.1004 | -7.5931
3.4916 | -34.1006 | -7.6046
3.4977 | -34.1007 | -7.6172
3.5038 | -34.1009 | -7.6289
3.5099 | -34.101 | -7.6402
3.5161 | -34.1013 | -7.6515
3.5222 | -34.1015 | -7.6633
3.5283 | -34.1015 | -7.675
3.5344 | -34.1016 | -7.6875
3.5405 | -34.1018 | -7.6991
3.5466 | -34.102 | -7.7106
3.5527 | -34.1022 | -7.7229
3.5588 | -34.1024 | -7.7344
3.5649 | -34.1025 | -7.7463
3.571 | -34.1027 | -7.7584
3.5771 | -34.1028 | -7.7709
3.5832 | -34.103 | -7.7821
3.5893 | -34.1033 | -7.7936
3.5954 | -34.1035 | -7.8058
3.6015 | -34.1037 | -7.8179
3.6076 | -34.104 | -7.8293
3.6137 | -34.1042 | -7.8409
3.6198 | -34.1046 | -7.8518
3.6259 | -34.1047 | -7.8643
3.632 | -34.105 | -7.8766
3.6381 | -34.1053 | -7.8881
3.6442 | -34.1056 | -7.9002
3.6503 | -34.1059 | -7.911
3.6565 | -34.1063 | -7.9243
3.6626 | -34.1064 | -7.9368
3.6687 | -34.1068 | -7.9491
3.6748 | -34.1072 | -7.9608
3.6809 | -34.1076 | -7.973
3.687 | -34.1079 | -7.986
3.6931 | -34.1083 | -7.9977
3.6992 | -34.1086 | -8.0099
3.7053 | -34.1089 | -8.0217
3.7114 | -34.1091 | -8.0345
3.7175 | -34.1096 | -8.0463
3.7236 | -34.1098 | -8.057
3.7297 | -34.11 | -8.0694
3.7358 | -34.1104 | -8.0827
3.7419 | -34.1107 | -8.0958
3.748 | -34.1109 | -8.1081
3.7541 | -34.1111 | -8.1208
3.7602 | -34.1115 | -8.1333
3.7663 | -34.1118 | -8.1453
3.7724 | -34.1122 | -8.158
3.7785 | -34.1126 | -8.1702
3.7846 | -34.1129 | -8.1824
3.7907 | -34.1134 | -8.1951
3.7969 | -34.1137 | -8.2084
3.803 | -34.1141 | -8.2203
3.8091 | -34.1145 | -8.2322
3.8152 | -34.1148 | -8.245
3.8213 | -34.1151 | -8.2583
3.8274 | -34.1156 | -8.2706
3.8335 | -34.116 | -8.2839
3.8396 | -34.1165 | -8.297
3.8457 | -34.117 | -8.3091
3.8518 | -34.1175 | -8.3218
3.8579 | -34.118 | -8.3341
3.864 | -34.1184 | -8.3467
3.8701 | -34.1189 | -8.3593
3.8762 | -34.1193 | -8.3722
3.8823 | -34.1196 | -8.3846
3.8884 | -34.1201 | -8.3967
3.8945 | -34.1207 | -8.4098
3.9006 | -34.1211 | -8.4231
3.9067 | -34.1216 | -8.4367
3.9128 | -34.122 | -8.4497
3.9189 | -34.1225 | -8.4622
3.925 | -34.1228 | -8.4748
3.9311 | -34.1233 | -8.4867
3.9372 | -34.1238 | -8.5
3.9434 | -34.1242 | -8.5123
3.9495 | -34.1247 | -8.5257
3.9556 | -34.1252 | -8.5387
3.9617 | -34.1256 | -8.5513
3.9678 | -34.1261 | -8.5639
3.9739 | -34.1266 | -8.5769
3.98 | -34.127 | -8.5905
3.9861 | -34.1274 | -8.6028
3.9922 | -34.1277 | -8.6162
3.9983 | -34.1282 | -8.6282
4.0044 | -34.1286 | -8.6398
4.0105 | -34.1291 | -8.6526
4.0166 | -34.1297 | -8.6654
4.0227 | -34.1302 | -8.6777
4.0288 | -34.1305 | -8.6909
4.0349 | -34.1309 | -8.7026
4.041 | -34.1313 | -8.7155
4.0471 | -34.1317 | -8.7277
4.0532 | -34.1322 | -8.7407
4.0593 | -34.1326 | -8.754
4.0654 | -34.133 | -8.7666
4.0715 | -34.1335 | -8.7798
4.0776 | -34.1339 | -8.7922
4.0838 | -34.1344 | -8.8055
4.0899 | -34.1348 | -8.8186
4.096 | -34.1352 | -8.8309
4.1021 | -34.1356 | -8.8435
4.1082 | -34.1361 | -8.8557
4.1143 | -34.1366 | -8.8676
4.1204 | -34.1371 | -8.8802
4.1265 | -34.1376 | -8.8927
4.1326 | -34.1381 | -8.9049
4.1387 | -34.1386 | -8.9179
4.1448 | -34.139 | -8.9307
4.1509 | -34.1394 | -8.9434
4.157 | -34.1398 | -8.9566
4.1631 | -34.1402 | -8.9697
4.1692 | -34.1406 | -8.9822
4.1753 | -34.1411 | -8.9949
4.1814 | -34.1415 | -9.0067
4.1875 | -34.1419 | -9.0199
4.1936 | -34.1423 | -9.033
4.1997 | -34.1428 | -9.0452
4.2058 | -34.1431 | -9.0586
4.2119 | -34.1435 | -9.0704
4.218 | -34.1439 | -9.0825
4.2241 | -34.1443 | -9.0942
4.2303 | -34.1449 | -9.1074
4.2364 | -34.1453 | -9.1207
4.2425 | -34.1457 | -9.1331
4.2486 | -34.1461 | -9.1461
4.2547 | -34.1463 | -9.1574
4.2608 | -34.1467 | -9.1694
4.2669 | -34.1471 | -9.1817
4.273 | -34.1475 | -9.1944
4.2791 | -34.1479 | -9.207
4.2852 | -34.1482 | -9.2193
4.2913 | -34.1485 | -9.2306
4.2974 | -34.1488 | -9.2435
4.3035 | -34.1492 | -9.2563
4.3096 | -34.1495 | -9.2687
4.3157 | -34.1498 | -9.2806
4.3218 | -34.1502 | -9.2926
4.3279 | -34.1506 | -9.305
4.334 | -34.1509 | -9.3168
4.3401 | -34.1512 | -9.3287
4.3462 | -34.1517 | -9.3415
4.3523 | -34.1521 | -9.3532
4.3584 | -34.1524 | -9.3638
4.3645 | -34.1527 | -9.3755
4.3707 | -34.153 | -9.3869
4.3768 | -34.1534 | -9.3992
4.3829 | -34.1536 | -9.4115
4.389 | -34.154 | -9.4233
4.3951 | -34.1544 | -9.4348
4.4012 | -34.1547 | -9.4467
4.4073 | -34.1551 | -9.4588
4.4134 | -34.1554 | -9.4709
4.4195 | -34.1558 | -9.4826
4.4256 | -34.1562 | -9.4942
4.4317 | -34.1565 | -9.5058
4.4378 | -34.1568 | -9.5173
4.4439 | -34.1571 | -9.5289
4.45 | -34.1574 | -9.5408
4.4561 | -34.1576 | -9.5526
4.4622 | -34.1579 | -9.5642
4.4683 | -34.1582 | -9.5755
4.4744 | -34.1585 | -9.5877
4.4805 | -34.1589 | -9.5988
4.4866 | -34.1592 | -9.6102
4.4927 | -34.1595 | -9.6224
4.4988 | -34.1598 | -9.6338
4.5049 | -34.1602 | -9.6449
4.511 | -34.1606 | -9.6559
4.5172 | -34.161 | -9.6678
4.5233 | -34.1613 | -9.6802
4.5294 | -34.1616 | -9.6917
4.5355 | -34.162 | -9.7031
4.5416 | -34.1623 | -9.7143
4.5477 | -34.1626 | -9.7256
4.5538 | -34.1629 | -9.7364
4.5599 | -34.1634 | -9.748
4.566 | -34.1637 | -9.7597
4.5721 | -34.1641 | -9.771
4.5782 | -34.1645 | -9.7828
4.5843 | -34.1648 | -9.7944
4.5904 | -34.1652 | -9.8058
4.5965 | -34.1656 | -9.8167
4.6026 | -34.166 | -9.8293
4.6087 | -34.1663 | -9.8404
4.6148 | -34.1667 | -9.8529
4.6209 | -34.1671 | -9.8636
4.627 | -34.1675 | -9.8756
4.6331 | -34.1677 | -9.8865
4.6392 | -34.1681 | -9.8978
4.6453 | -34.1685 | -9.9092
4.6514 | -34.1689 | -9.9207
4.6576 | -34.169 | -9.9331
4.6637 | -34.1694 | -9.9429
4.6698 | -34.1697 | -9.9535
4.6759 | -34.17 | -9.9653
4.682 | -34.1703 | -9.9765
4.6881 | -34.1706 | -9.9874
4.6942 | -34.1709 | -9.9988
4.7003 | -34.1713 | -10.0103
4.7064 | -34.1716 | -10.0215
4.7125 | -34.1721 | -10.033
4.7186 | -34.1723 | -10.0436
4.7247 | -34.1727 | -10.0552
4.7308 | -34.1732 | -10.0666
4.7369 | -34.1735 | -10.0782
4.743 | -34.1738 | -10.0905
4.7491 | -34.1741 | -10.1021
4.7552 | -34.1743 | -10.1134
4.7613 | -34.1746 | -10.1249
4.7674 | -34.175 | -10.1356
4.7735 | -34.1752 | -10.1477
4.7796 | -34.1756 | -10.1595
4.7857 | -34.1758 | -10.1711
4.7918 | -34.1761 | -10.1818
4.7979 | -34.1766 | -10.193
4.8041 | -34.1768 | -10.2043
4.8102 | -34.1772 | -10.2153
4.8163 | -34.1776 | -10.2268
4.8224 | -34.1778 | -10.2376
4.8285 | -34.1782 | -10.2486
4.8346 | -34.1786 | -10.2601
4.8407 | -34.1789 | -10.2715
4.8468 | -34.1792 | -10.2831
4.8529 | -34.1796 | -10.2949
4.859 | -34.1799 | -10.3064
4.8651 | -34.1802 | -10.3177
4.8712 | -34.1805 | -10.3291
4.8773 | -34.1809 | -10.3407
4.8834 | -34.1812 | -10.3523
4.8895 | -34.1816 | -10.3642
4.8956 | -34.1819 | -10.3761
4.9017 | -34.1821 | -10.3876
4.9078 | -34.1825 | -10.3994
4.9139 | -34.1828 | -10.41
4.92 | -34.1832 | -10.4211
4.9261 | -34.1836 | -10.4319
4.9322 | -34.1841 | -10.4431
4.9383 | -34.1845 | -10.4541
4.9445 | -34.1848 | -10.4662
4.9506 | -34.1852 | -10.477
4.9567 | -34.1855 | -10.4883
4.9628 | -34.1859 | -10.4996
4.9689 | -34.1862 | -10.5114
4.975 | -34.1865 | -10.5227
4.9811 | -34.1869 | -10.5345
4.9872 | -34.1873 | -10.5462
4.9933 | -34.1877 | -10.5578
4.9994 | -34.1882 | -10.5697
5.0055 | -34.1886 | -10.5813
5.0116 | -34.189 | -10.5934
5.0177 | -34.1895 | -10.6054
5.0238 | -34.1899 | -10.6178
5.0299 | -34.1905 | -10.6306
5.036 | -34.1908 | -10.6421
5.0421 | -34.1911 | -10.6538
5.0482 | -34.1915 | -10.666
5.0543 | -34.1919 | -10.6772
5.0604 | -34.1924 | -10.6894
5.0665 | -34.1929 | -10.701
5.0726 | -34.1933 | -10.7133
5.0787 | -34.1937 | -10.7254
5.0848 | -34.1941 | -10.7372
5.091 | -34.1945 | -10.749
5.0971 | -34.1949 | -10.7611
5.1032 | -34.1954 | -10.7734
5.1093 | -34.1958 | -10.7852
5.1154 | -34.1961 | -10.7971
5.1215 | -34.1965 | -10.809
5.1276 | -34.1969 | -10.8207
5.1337 | -34.1975 | -10.8323
5.1398 | -34.1978 | -10.8441
5.1459 | -34.1983 | -10.8563
5.152 | -34.1987 | -10.8692
5.1581 | -34.199 | -10.8817
5.1642 | -34.1994 | -10.8936
5.1703 | -34.1999 | -10.9055
5.1764 | -34.2003 | -10.9173
5.1825 | -34.2007 | -10.9295
5.1886 | -34.2011 | -10.9412
5.1947 | -34.2015 | -10.9538
5.2008 | -34.202 | -10.9656
5.2069 | -34.2025 | -10.9785
5.213 | -34.2029 | -10.99
5.2191 | -34.2034 | -11.0023
5.2252 | -34.2038 | -11.0144
5.2314 | -34.2041 | -11.027
5.2375 | -34.2046 | -11.0388
5.2436 | -34.205 | -11.0509
5.2497 | -34.2055 | -11.0626
5.2558 | -34.2059 | -11.0754
5.2619 | -34.2063 | -11.0882
5.268 | -34.2068 | -11.1005
5.2741 | -34.2073 | -11.1135
5.2802 | -34.2077 | -11.1261
5.2863 | -34.2082 | -11.1382
5.2924 | -34.2086 | -11.1512
5.2985 | -34.2091 | -11.1643
5.3046 | -34.2097 | -11.1773
5.3107 | -34.2101 | -11.19
5.3168 | -34.2107 | -11.2023
5.3229 | -34.2111 | -11.214
5.329 | -34.2116 | -11.2259
5.3351 | -34.2121 | -11.2379
5.3412 | -34.2126 | -11.25
5.3473 | -34.2131 | -11.2627
5.3534 | -34.2136 | -11.2754
5.3595 | -34.214 | -11.288
5.3656 | -34.2145 | -11.3008
5.3717 | -34.2149 | -11.3131
5.3779 | -34.2154 | -11.3259
5.384 | -34.2158 | -11.3384
5.3901 | -34.2163 | -11.3509
5.3962 | -34.2167 | -11.3636
5.4023 | -34.2172 | -11.3761
5.4084 | -34.2177 | -11.3884
5.4145 | -34.2181 | -11.4012
5.4206 | -34.2185 | -11.4143
5.4267 | -34.219 | -11.4275
5.4328 | -34.2195 | -11.4395
5.4389 | -34.22 | -11.4516
5.445 | -34.2203 | -11.4647
5.4511 | -34.2209 | -11.4769
5.4572 | -34.2214 | -11.4899
5.4633 | -34.2218 | -11.5022
5.4694 | -34.2222 | -11.5145
5.4755 | -34.2227 | -11.5272
5.4816 | -34.2231 | -11.5397
5.4877 | -34.2236 | -11.5521
5.4938 | -34.2241 | -11.5644
5.4999 | -34.2245 | -11.5768
5.506 | -34.225 | -11.5894
5.5121 | -34.2255 | -11.6023
5.5183 | -34.2259 | -11.6142
5.5244 | -34.2263 | -11.6269
5.5305 | -34.2268 | -11.6395
5.5366 | -34.2273 | -11.6518
5.5427 | -34.2277 | -11.6649
5.5488 | -34.2281 | -11.6777
5.5549 | -34.2285 | -11.6901
5.561 | -34.229 | -11.7032
5.5671 | -34.2295 | -11.7154
5.5732 | -34.2299 | -11.7281
5.5793 | -34.2304 | -11.7414
5.5854 | -34.231 | -11.754
5.5915 | -34.2313 | -11.7668
5.5976 | -34.2317 | -11.7796
5.6037 | -34.2322 | -11.7921
5.6098 | -34.2326 | -11.8044
5.6159 | -34.2331 | -11.8172
5.622 | -34.2336 | -11.8294
5.6281 | -34.234 | -11.8421
5.6342 | -34.2345 | -11.8548
5.6403 | -34.2349 | -11.8669
5.6464 | -34.2354 | -11.8795
5.6525 | -34.2358 | -11.8918
5.6586 | -34.2363 | -11.9036
5.6648 | -34.2367 | -11.9167
5.6709 | -34.2372 | -11.9296
5.677 | -34.2377 | -11.9428
5.6831 | -34.2381 | -11.9549
5.6892 | -34.2386 | -11.9675
5.6953 | -34.239 | -11.9799
5.7014 | -34.2395 | -11.9929
5.7075 | -34.2399 | -12.0057
5.7136 | -34.2403 | -12.018
5.7197 | -34.2408 | -12.0305
5.7258 | -34.2412 | -12.0428
5.7319 | -34.2416 | -12.0553
5.738 | -34.2421 | -12.0678
5.7441 | -34.2426 | -12.0803
5.7502 | -34.2431 | -12.0924
5.7563 | -34.2435 | -12.1047
5.7624 | -34.244 | -12.1179
5.7685 | -34.2444 | -12.1302
5.7746 | -34.2447 | -12.1425
5.7807 | -34.2452 | -12.1547
5.7868 | -34.2456 | -12.1672
5.7929 | -34.246 | -12.1795
5.799 | -34.2465 | -12.1918
5.8052 | -34.2469 | -12.2044
5.8113 | -34.2472 | -12.2168
5.8174 | -34.2476 | -12.2294
5.8235 | -34.2481 | -12.2416
5.8296 | -34.2485 | -12.2546
5.8357 | -34.2489 | -12.267
5.8418 | -34.2493 | -12.2797
5.8479 | -34.2497 | -12.2924
5.854 | -34.2501 | -12.3051
5.8601 | -34.2505 | -12.3174
5.8662 | -34.2509 | -12.3301
5.8723 | -34.2513 | -12.3419
5.8784 | -34.2518 | -12.3543
5.8845 | -34.2522 | -12.3669
5.8906 | -34.2526 | -12.3793
5.8967 | -34.2531 | -12.3913
5.9028 | -34.2535 | -12.4038
5.9089 | -34.2539 | -12.4156
5.915 | -34.2543 | -12.4279
5.9211 | -34.2548 | -12.4403
5.9272 | -34.2552 | -12.4526
5.9333 | -34.2556 | -12.4644
5.9394 | -34.2561 | -12.4766
5.9455 | -34.2565 | -12.4885
5.9517 | -34.2571 | -12.5009
5.9578 | -34.2575 | -12.5135
5.9639 | -34.258 | -12.5255
5.97 | -34.2584 | -12.538
5.9761 | -34.2588 | -12.5499
5.9822 | -34.2593 | -12.5624
5.9883 | -34.2598 | -12.5745
5.9944 | -34.2602 | -12.5869
6.0005 | -34.2607 | -12.5994
6.0066 | -34.2611 | -12.6113
6.0127 | -34.2616 | -12.6229
6.0188 | -34.262 | -12.6356
6.0249 | -34.2624 | -12.6478
6.031 | -34.2629 | -12.6598
6.0371 | -34.2634 | -12.6723
6.0432 | -34.2638 | -12.6846
6.0493 | -34.2643 | -12.6966
6.0554 | -34.2648 | -12.7088
6.0615 | -34.2653 | -12.7215
6.0676 | -34.2657 | -12.7343
6.0737 | -34.2662 | -12.7467
6.0798 | -34.2665 | -12.7588
6.0859 | -34.267 | -12.771
6.0921 | -34.2675 | -12.783
6.0982 | -34.268 | -12.7954
6.1043 | -34.2685 | -12.8075
6.1104 | -34.269 | -12.8203
6.1165 | -34.2695 | -12.8317
6.1226 | -34.2699 | -12.8434
6.1287 | -34.2704 | -12.8557
6.1348 | -34.271 | -12.8679
6.1409 | -34.2714 | -12.88
6.147 | -34.2719 | -12.8921
6.1531 | -34.2723 | -12.9042
6.1592 | -34.2728 | -12.9164
6.1653 | -34.2732 | -12.9283
6.1714 | -34.2737 | -12.9406
6.1775 | -34.2742 | -12.953
6.1836 | -34.2746 | -12.965
6.1897 | -34.2751 | -12.9772
6.1958 | -34.2755 | -12.9895
6.2019 | -34.276 | -13.0019
6.208 | -34.2765 | -13.0145
6.2141 | -34.277 | -13.0268
6.2202 | -34.2774 | -13.0394
6.2263 | -34.2779 | -13.0514
6.2325 | -34.2784 | -13.0631
6.2386 | -34.2789 | -13.0749
6.2447 | -34.2794 | -13.0873
6.2508 | -34.2798 | -13.0993
6.2569 | -34.2803 | -13.1117
6.263 | -34.2808 | -13.1236
6.2691 | -34.2813 | -13.1357
6.2752 | -34.2818 | -13.1477
6.2813 | -34.2823 | -13.1598
6.2874 | -34.2827 | -13.172
6.2935 | -34.2832 | -13.1838
6.2996 | -34.2837 | -13.196
6.3057 | -34.2842 | -13.2081
6.3118 | -34.2847 | -13.2201
6.3179 | -34.2851 | -13.2322
6.324 | -34.2856 | -13.2446
6.3301 | -34.2861 | -13.2564
6.3362 | -34.2866 | -13.2686
6.3423 | -34.2872 | -13.2808
6.3484 | -34.2877 | -13.2931
6.3545 | -34.2882 | -13.3047
6.3606 | -34.2887 | -13.3165
6.3667 | -34.2892 | -13.3283
6.3728 | -34.2898 | -13.3405
6.379 | -34.2902 | -13.3525
6.3851 | -34.2907 | -13.3646
6.3912 | -34.2912 | -13.3765
6.3973 | -34.2917 | -13.3882
6.4034 | -34.2922 | -13.4002
6.4095 | -34.2927 | -13.4125
6.4156 | -34.2931 | -13.4243
6.4217 | -34.2936 | -13.4363
6.4278 | -34.2941 | -13.4482
6.4339 | -34.2947 | -13.4602
6.44 | -34.2952 | -13.4719
6.4461 | -34.2956 | -13.4844
6.4522 | -34.2961 | -13.4958
6.4583 | -34.2966 | -13.5079
6.4644 | -34.297 | -13.5203
6.4705 | -34.2976 | -13.5324
6.4766 | -34.298 | -13.5448
6.4827 | -34.2985 | -13.5572
6.4888 | -34.299 | -13.5691
6.4949 | -34.2995 | -13.5807
6.501 | -34.3 | -13.593
6.5071 | -34.3005 | -13.6051
6.5132 | -34.301 | -13.617
6.5194 | -34.3015 | -13.6289
6.5255 | -34.302 | -13.6406
6.5316 | -34.3025 | -13.6527
6.5377 | -34.303 | -13.665
6.5438 | -34.3035 | -13.6768
6.5499 | -34.304 | -13.6886
6.556 | -34.3045 | -13.7004
6.5621 | -34.305 | -13.7126
6.5682 | -34.3054 | -13.7245
6.5743 | -34.3059 | -13.7371
6.5804 | -34.3064 | -13.7492
6.5865 | -34.3068 | -13.7614
6.5926 | -34.3074 | -13.7733
6.5987 | -34.3079 | -13.785
6.6048 | -34.3083 | -13.7976
6.6109 | -34.3089 | -13.8095
6.617 | -34.3093 | -13.8224
6.6231 | -34.3097 | -13.8346
6.6292 | -34.3103 | -13.8465
6.6353 | -34.3108 | -13.8594
6.6414 | -34.3114 | -13.8717
6.6475 | -34.3119 | -13.8838
6.6536 | -34.3123 | -13.8955
6.6597 | -34.3129 | -13.9073
6.6659 | -34.3134 | -13.9192
6.672 | -34.3139 | -13.9312
6.6781 | -34.3144 | -13.943
6.6842 | -34.3149 | -13.9552
6.6903 | -34.3154 | -13.9671
6.6964 | -34.3159 | -13.9792
6.7025 | -34.3164 | -13.9911
6.7086 | -34.3169 | -14.0029
6.7147 | -34.3174 | -14.015
6.7208 | -34.3179 | -14.0275
6.7269 | -34.3184 | -14.0391
6.733 | -34.3188 | -14.0504
6.7391 | -34.3194 | -14.0625
6.7452 | -34.3199 | -14.074
6.7513 | -34.3204 | -14.0856
6.7574 | -34.321 | -14.0978
6.7635 | -34.3215 | -14.1093
6.7696 | -34.3221 | -14.1213
6.7757 | -34.3226 | -14.1336
6.7818 | -34.3231 | -14.1456
6.7879 | -34.3236 | -14.1575
6.794 | -34.3241 | -14.1694
6.8001 | -34.3246 | -14.1811
6.8063 | -34.3251 | -14.1932
6.8124 | -34.3256 | -14.2053
6.8185 | -34.3261 | -14.2173
6.8246 | -34.3266 | -14.2292
6.8307 | -34.3271 | -14.241
6.8368 | -34.3276 | -14.2532
6.8429 | -34.3281 | -14.2651
6.849 | -34.3286 | -14.277
6.8551 | -34.3292 | -14.2889
6.8612 | -34.3297 | -14.3007
6.8673 | -34.3302 | -14.3124
6.8734 | -34.3307 | -14.325
6.8795 | -34.3312 | -14.3372
6.8856 | -34.3317 | -14.3486
6.8917 | -34.3322 | -14.36
6.8978 | -34.3327 | -14.3719
6.9039 | -34.3332 | -14.3836
6.91 | -34.3338 | -14.3956
6.9161 | -34.3342 | -14.4073
6.9222 | -34.3347 | -14.4192
6.9283 | -34.3352 | -14.4307
6.9344 | -34.3357 | -14.4424
6.9405 | -34.3362 | -14.4541
6.9466 | -34.3367 | -14.4663
6.9528 | -34.3372 | -14.4782
6.9589 | -34.3377 | -14.49
6.965 | -34.3383 | -14.5018
6.9711 | -34.3388 | -14.5136
6.9772 | -34.3394 | -14.5253
6.9833 | -34.3398 | -14.5373
6.9894 | -34.3403 | -14.5496
6.9955 | -34.3408 | -14.5613
7.0016 | -34.3413 | -14.5735
7.0077 | -34.3418 | -14.5856
7.0138 | -34.3424 | -14.5978
7.0199 | -34.3428 | -14.6094
7.026 | -34.3432 | -14.6211
7.0321 | -34.3437 | -14.6329
7.0382 | -34.3442 | -14.6445
7.0443 | -34.3448 | -14.6566
7.0504 | -34.3452 | -14.6687
7.0565 | -34.3457 | -14.6807
7.0626 | -34.3462 | -14.6924
7.0687 | -34.3467 | -14.7037
7.0748 | -34.3473 | -14.7154
7.0809 | -34.3477 | -14.7273
7.087 | -34.3482 | -14.7392
7.0932 | -34.3487 | -14.7509
7.0993 | -34.3492 | -14.7628
7.1054 | -34.3497 | -14.7743
7.1115 | -34.3501 | -14.786
7.1176 | -34.3506 | -14.7976
7.1237 | -34.351 | -14.8094
7.1298 | -34.3514 | -14.8208
7.1359 | -34.3519 | -14.8322
7.142 | -34.3524 | -14.844
7.1481 | -34.3529 | -14.8555
7.1542 | -34.3533 | -14.8671
7.1603 | -34.3538 | -14.8787
7.1664 | -34.3543 | -14.8905
7.1725 | -34.3548 | -14.9019
7.1786 | -34.3554 | -14.9137
7.1847 | -34.3558 | -14.9253
7.1908 | -34.3563 | -14.9372
7.1969 | -34.3568 | -14.9488
7.203 | -34.3573 | -14.9605
7.2091 | -34.3578 | -14.9719
7.2152 | -34.3582 | -14.9836
7.2213 | -34.3587 | -14.995
7.2274 | -34.3592 | -15.0067
7.2335 | -34.3597 | -15.0183
7.2397 | -34.3602 | -15.03
7.2458 | -34.3606 | -15.0411
7.2519 | -34.3611 | -15.0527
7.258 | -34.3616 | -15.0645
7.2641 | -34.3621 | -15.0761
7.2702 | -34.3626 | -15.088
7.2763 | -34.3631 | -15.0994
7.2824 | -34.3637 | -15.1109
7.2885 | -34.3641 | -15.1225
7.2946 | -34.3646 | -15.1342
7.3007 | -34.365 | -15.1459
7.3068 | -34.3655 | -15.1577
7.3129 | -34.366 | -15.1691
7.319 | -34.3665 | -15.1808
7.3251 | -34.367 | -15.192
7.3312 | -34.3675 | -15.2033
7.3373 | -34.368 | -15.2146
7.3434 | -34.3686 | -15.2259
7.3495 | -34.3691 | -15.237
7.3556 | -34.3695 | -15.2489
7.3617 | -34.37 | -15.2609
7.3678 | -34.3705 | -15.2725
7.3739 | -34.371 | -15.2837
7.3801 | -34.3715 | -15.2952
7.3862 | -34.3719 | -15.3066
7.3923 | -34.3724 | -15.3181
7.3984 | -34.3729 | -15.3295
7.4045 | -34.3734 | -15.3409
7.4106 | -34.3739 | -15.3523
7.4167 | -34.3744 | -15.364
7.4228 | -34.3749 | -15.3754
7.4289 | -34.3754 | -15.3866
7.435 | -34.3759 | -15.3979
7.4411 | -34.3764 | -15.4091
7.4472 | -34.3769 | -15.4208
7.4533 | -34.3774 | -15.4321
7.4594 | -34.3779 | -15.4439
7.4655 | -34.3784 | -15.4553
7.4716 | -34.3789 | -15.467
7.4777 | -34.3794 | -15.478
7.4838 | -34.3799 | -15.4896
7.4899 | -34.3804 | -15.501
7.496 | -34.3809 | -15.5125
7.5021 | -34.3814 | -15.5238
7.5082 | -34.3819 | -15.5353
7.5143 | -34.3824 | -15.5467
7.5204 | -34.3829 | -15.5579
7.5266 | -34.3833 | -15.5693
7.5327 | -34.3838 | -15.5806
7.5388 | -34.3843 | -15.592
7.5449 | -34.3847 | -15.6036
7.551 | -34.3852 | -15.6152
7.5571 | -34.3857 | -15.6265
7.5632 | -34.3862 | -15.6378
7.5693 | -34.3867 | -15.6493
7.5754 | -34.3873 | -15.6608
7.5815 | -34.3878 | -15.6718
7.5876 | -34.3882 | -15.6832
7.5937 | -34.3887 | -15.6945
7.5998 | -34.3892 | -15.7059
7.6059 | -34.3896 | -15.7173
7.612 | -34.3901 | -15.7286
7.6181 | -34.3906 | -15.7399
7.6242 | -34.3911 | -15.7514
7.6303 | -34.3917 | -15.7626
7.6364 | -34.3921 | -15.7742
7.6425 | -34.3926 | -15.7859
7.6486 | -34.3931 | -15.7973
7.6547 | -34.3936 | -15.8088
7.6608 | -34.3942 | -15.8201
7.667 | -34.3946 | -15.831
7.6731 | -34.3951 | -15.8425
7.6792 | -34.3955 | -15.854
7.6853 | -34.396 | -15.8653
7.6914 | -34.3964 | -15.8768
7.6975 | -34.3969 | -15.8884
7.7036 | -34.3974 | -15.8998
7.7097 | -34.3979 | -15.9111
7.7158 | -34.3984 | -15.9224
7.7219 | -34.3989 | -15.9336
7.728 | -34.3993 | -15.9451
7.7341 | -34.3998 | -15.9565
7.7402 | -34.4003 | -15.9678
7.7463 | -34.4008 | -15.9792
7.7524 | -34.4012 | -15.9908
7.7585 | -34.4017 | -16.0023
7.7646 | -34.4021 | -16.0134
7.7707 | -34.4026 | -16.0246
7.7768 | -34.403 | -16.0358
7.7829 | -34.4034 | -16.0472
7.789 | -34.4039 | -16.0584
7.7951 | -34.4044 | -16.0701
7.8012 | -34.4049 | -16.0818
7.8073 | -34.4053 | -16.093
7.8135 | -34.4058 | -16.1045
7.8196 | -34.4063 | -16.1158
7.8257 | -34.4068 | -16.1275
7.8318 | -34.4072 | -16.1392
7.8379 | -34.4077 | -16.1502
7.844 | -34.4081 | -16.1611
7.8501 | -34.4085 | -16.1727
7.8562 | -34.409 | -16.1846
7.8623 | -34.4095 | -16.196
7.8684 | -34.4099 | -16.2071
7.8745 | -34.4104 | -16.2184
7.8806 | -34.4108 | -16.2292
7.8867 | -34.4113 | -16.2404
7.8928 | -34.4118 | -16.2517
7.8989 | -34.4124 | -16.2626
7.905 | -34.4128 | -16.2744
7.9111 | -34.4133 | -16.2857
7.9172 | -34.4138 | -16.2968
7.9233 | -34.4142 | -16.3082
7.9294 | -34.4147 | -16.3197
7.9355 | -34.4152 | -16.3306
7.9416 | -34.4156 | -16.3419
7.9477 | -34.4161 | -16.3527
7.9539 | -34.4165 | -16.3637
7.96 | -34.417 | -16.3753
7.9661 | -34.4175 | -16.3867
7.9722 | -34.418 | -16.3981
7.9783 | -34.4184 | -16.4094
7.9844 | -34.4189 | -16.421
7.9905 | -34.4194 | -16.4322
7.9966 | -34.4198 | -16.4436
8.0027 | -34.4203 | -16.4551
8.0088 | -34.4207 | -16.4661
8.0149 | -34.4212 | -16.4771
8.021 | -34.4217 | -16.4885
8.0271 | -34.4221 | -16.5002
8.0332 | -34.4226 | -16.5114
8.0393 | -34.4231 | -16.5229
8.0454 | -34.4235 | -16.5344
8.0515 | -34.424 | -16.5456
8.0576 | -34.4244 | -16.5572
8.0637 | -34.4249 | -16.568
8.0698 | -34.4254 | -16.5789
8.0759 | -34.4258 | -16.5899
8.082 | -34.4263 | -16.6014
8.0881 | -34.4267 | -16.6125
8.0942 | -34.4272 | -16.6242
8.1004 | -34.4277 | -16.6359
8.1065 | -34.4282 | -16.6471
8.1126 | -34.4286 | -16.6583
8.1187 | -34.4291 | -16.6697
8.1248 | -34.4296 | -16.6811
8.1309 | -34.43 | -16.6924
8.137 | -34.4305 | -16.7036
8.1431 | -34.431 | -16.7151
8.1492 | -34.4315 | -16.7265
8.1553 | -34.4319 | -16.7378
8.1614 | -34.4324 | -16.7493
8.1675 | -34.4328 | -16.7605
8.1736 | -34.4333 | -16.772
8.1797 | -34.4337 | -16.7835
8.1858 | -34.4342 | -16.7948
8.1919 | -34.4346 | -16.8059
8.198 | -34.4351 | -16.8169
8.2041 | -34.4356 | -16.8285
8.2102 | -34.4361 | -16.8398
8.2163 | -34.4366 | -16.8511
8.2224 | -34.437 | -16.8626
8.2285 | -34.4375 | -16.8736
8.2346 | -34.438 | -16.8848
8.2408 | -34.4385 | -16.8963
8.2469 | -34.4389 | -16.9075
8.253 | -34.4394 | -16.9187
8.2591 | -34.4399 | -16.9297
8.2652 | -34.4404 | -16.9408
8.2713 | -34.4409 | -16.9522
8.2774 | -34.4414 | -16.9635
8.2835 | -34.4419 | -16.9752
8.2896 | -34.4423 | -16.986
8.2957 | -34.4428 | -16.9973
8.3018 | -34.4433 | -17.0085
8.3079 | -34.4438 | -17.0198
8.314 | -34.4442 | -17.0309
8.3201 | -34.4447 | -17.0421
8.3262 | -34.4452 | -17.0533
8.3323 | -34.4457 | -17.0645
8.3384 | -34.4462 | -17.0762
8.3445 | -34.4466 | -17.087
8.3506 | -34.4471 | -17.0985
8.3567 | -34.4476 | -17.11
8.3628 | -34.448 | -17.1212
8.3689 | -34.4485 | -17.1324
8.375 | -34.449 | -17.1437
8.3812 | -34.4495 | -17.1549
8.3873 | -34.4499 | -17.1663
8.3934 | -34.4504 | -17.1775
8.3995 | -34.4509 | -17.1888
8.4056 | -34.4514 | -17.2002
8.4117 | -34.4518 | -17.2117
8.4178 | -34.4523 | -17.2226
8.4239 | -34.4528 | -17.2336
8.43 | -34.4533 | -17.2449
8.4361 | -34.4538 | -17.2565
8.4422 | -34.4542 | -17.2677
8.4483 | -34.4547 | -17.2789
8.4544 | -34.4552 | -17.2904
8.4605 | -34.4556 | -17.3015
8.4666 | -34.4561 | -17.313
8.4727 | -34.4565 | -17.3242
8.4788 | -34.457 | -17.3352
8.4849 | -34.4575 | -17.3468
8.491 | -34.4579 | -17.3579
8.4971 | -34.4584 | -17.3692
8.5032 | -34.4589 | -17.3805
8.5093 | -34.4594 | -17.3917
8.5154 | -34.4599 | -17.4027
8.5215 | -34.4603 | -17.414
8.5277 | -34.4608 | -17.4254
8.5338 | -34.4612 | -17.4364
8.5399 | -34.4617 | -17.4479
8.546 | -34.4622 | -17.4589
8.5521 | -34.4627 | -17.4702
8.5582 | -34.4632 | -17.4816
8.5643 | -34.4636 | -17.4927
8.5704 | -34.4641 | -17.5036
8.5765 | -34.4646 | -17.5148
8.5826 | -34.4651 | -17.5259
8.5887 | -34.4656 | -17.5371
8.5948 | -34.466 | -17.5479
8.6009 | -34.4665 | -17.5591
8.607 | -34.467 | -17.5705
8.6131 | -34.4674 | -17.582
8.6192 | -34.4679 | -17.5929
8.6253 | -34.4684 | -17.6042
8.6314 | -34.4689 | -17.6154
8.6375 | -34.4694 | -17.6268
8.6436 | -34.4699 | -17.6381
8.6497 | -34.4704 | -17.6493
8.6558 | -34.4709 | -17.6603
8.6619 | -34.4713 | -17.6717
8.6681 | -34.4718 | -17.6828
8.6742 | -34.4723 | -17.6942
8.6803 | -34.4728 | -17.7056
8.6864 | -34.4733 | -17.717
8.6925 | -34.4738 | -17.7282
8.6986 | -34.4743 | -17.7398
8.7047 | -34.4748 | -17.7511
8.7108 | -34.4753 | -17.7624
8.7169 | -34.4758 | -17.7734
8.723 | -34.4763 | -17.7846
8.7291 | -34.4767 | -17.7957
8.7352 | -34.4773 | -17.8072
8.7413 | -34.4777 | -17.8183
8.7474 | -34.4782 | -17.8292
8.7535 | -34.4788 | -17.8406
8.7596 | -34.4793 | -17.8522
8.7657 | -34.4798 | -17.8634
8.7718 | -34.4802 | -17.8744
8.7779 | -34.4807 | -17.8861
8.784 | -34.4812 | -17.8972
8.7901 | -34.4816 | -17.9088
8.7962 | -34.4822 | -17.9201
8.8023 | -34.4827 | -17.931
8.8084 | -34.4832 | -17.9424
8.8146 | -34.4836 | -17.9535
8.8207 | -34.4842 | -17.9647
8.8268 | -34.4847 | -17.9758
8.8329 | -34.4852 | -17.9871
8.839 | -34.4857 | -17.9984
8.8451 | -34.4862 | -18.0095
8.8512 | -34.4867 | -18.0207
8.8573 | -34.4872 | -18.0318
8.8634 | -34.4878 | -18.0429
8.8695 | -34.4882 | -18.0543
8.8756 | -34.4887 | -18.0653
8.8817 | -34.4892 | -18.0767
8.8878 | -34.4897 | -18.0881
8.8939 | -34.4901 | -18.0991
8.9 | -34.4907 | -18.11
8.9061 | -34.4912 | -18.1212
8.9122 | -34.4917 | -18.1323
8.9183 | -34.4922 | -18.1434
8.9244 | -34.4927 | -18.1545
8.9305 | -34.4932 | -18.1656
8.9366 | -34.4937 | -18.1767
8.9427 | -34.4943 | -18.188
8.9488 | -34.4948 | -18.1991
8.955 | -34.4953 | -18.2103
8.9611 | -34.4958 | -18.2216
8.9672 | -34.4963 | -18.2331
8.9733 | -34.4968 | -18.2441
8.9794 | -34.4973 | -18.2554
8.9855 | -34.4979 | -18.2666
8.9916 | -34.4983 | -18.2776
8.9977 | -34.4989 | -18.2887
9.0038 | -34.4994 | -18.3
9.0099 | -34.5 | -18.3113
9.016 | -34.5005 | -18.3225
9.0221 | -34.501 | -18.3338
9.0282 | -34.5015 | -18.3451
9.0343 | -34.5021 | -18.3565
9.0404 | -34.5026 | -18.3675
9.0465 | -34.5032 | -18.379
9.0526 | -34.5037 | -18.3901
9.0587 | -34.5042 | -18.4012
9.0648 | -34.5048 | -18.4123
9.0709 | -34.5053 | -18.4233
9.077 | -34.5058 | -18.435
9.0831 | -34.5063 | -18.4462
9.0892 | -34.5068 | -18.4573
9.0953 | -34.5073 | -18.4686
9.1015 | -34.5078 | -18.4798
9.1076 | -34.5084 | -18.491
9.1137 | -34.5089 | -18.5021
9.1198 | -34.5095 | -18.5129
9.1259 | -34.51 | -18.5242
9.132 | -34.5106 | -18.5356
9.1381 | -34.5111 | -18.5467
9.1442 | -34.5117 | -18.5574
9.1503 | -34.5122 | -18.5687
9.1564 | -34.5127 | -18.5799
9.1625 | -34.5133 | -18.5911
9.1686 | -34.5138 | -18.6023
9.1747 | -34.5143 | -18.6135
9.1808 | -34.5149 | -18.6248
9.1869 | -34.5154 | -18.6363
9.193 | -34.5159 | -18.6476
9.1991 | -34.5164 | -18.6585
9.2052 | -34.517 | -18.6696
9.2113 | -34.5176 | -18.681
9.2174 | -34.5181 | -18.6926
9.2235 | -34.5187 | -18.7035
9.2296 | -34.5193 | -18.7148
9.2357 | -34.5198 | -18.726
9.2419 | -34.5203 | -18.7371
9.248 | -34.5209 | -18.7485
9.2541 | -34.5214 | -18.7598
9.2602 | -34.5219 | -18.7711
9.2663 | -34.5225 | -18.7821
9.2724 | -34.523 | -18.7933
9.2785 | -34.5236 | -18.8045
9.2846 | -34.5241 | -18.8157
9.2907 | -34.5247 | -18.8268
9.2968 | -34.5252 | -18.8381
9.3029 | -34.5257 | -18.8494
9.309 | -34.5263 | -18.8605
9.3151 | -34.5268 | -18.8716
9.3212 | -34.5274 | -18.8828
9.3273 | -34.528 | -18.8939
9.3334 | -34.5286 | -18.9049
9.3395 | -34.5292 | -18.9159
9.3456 | -34.5297 | -18.9269
9.3517 | -34.5303 | -18.9377
9.3578 | -34.5309 | -18.9487
9.3639 | -34.5314 | -18.96
9.37 | -34.532 | -18.9712
9.3761 | -34.5325 | -18.9823
9.3822 | -34.5331 | -18.9933
9.3884 | -34.5337 | -19.0044
9.3945 | -34.5342 | -19.0156
9.4006 | -34.5348 | -19.0268
9.4067 | -34.5353 | -19.0381
9.4128 | -34.5359 | -19.0491
9.4189 | -34.5365 | -19.0604
9.425 | -34.5371 | -19.0714
9.4311 | -34.5376 | -19.0825
9.4372 | -34.5382 | -19.0936
9.4433 | -34.5387 | -19.1049
9.4494 | -34.5393 | -19.116
9.4555 | -34.5398 | -19.1271
9.4616 | -34.5404 | -19.1381
9.4677 | -34.541 | -19.1491
9.4738 | -34.5416 | -19.1605
9.4799 | -34.5421 | -19.1716
9.486 | -34.5426 | -19.183
9.4921 | -34.5432 | -19.1945
9.4982 | -34.5438 | -19.2052
9.5043 | -34.5443 | -19.2161
9.5104 | -34.5449 | -19.2271
9.5165 | -34.5455 | -19.2382
9.5226 | -34.5461 | -19.2495
9.5288 | -34.5467 | -19.2605
9.5349 | -34.5472 | -19.2716
9.541 | -34.5478 | -19.2827
9.5471 | -34.5483 | -19.2939
9.5532 | -34.5489 | -19.305
9.5593 | -34.5495 | -19.3161
9.5654 | -34.5501 | -19.3272
9.5715 | -34.5507 | -19.3386
9.5776 | -34.5512 | -19.3499
9.5837 | -34.5518 | -19.3608
9.5898 | -34.5523 | -19.3721
9.5959 | -34.5529 | -19.3831
9.602 | -34.5534 | -19.3942
9.6081 | -34.554 | -19.4054
9.6142 | -34.5546 | -19.4164
9.6203 | -34.5551 | -19.4276
9.6264 | -34.5556 | -19.4383
9.6325 | -34.5561 | -19.4496
9.6386 | -34.5567 | -19.4606
9.6447 | -34.5573 | -19.4716
9.6508 | -34.5579 | -19.4828
9.6569 | -34.5584 | -19.4938
9.663 | -34.559 | -19.505
9.6691 | -34.5597 | -19.516
9.6753 | -34.5603 | -19.5271
9.6814 | -34.5609 | -19.538
9.6875 | -34.5615 | -19.5489
9.6936 | -34.562 | -19.5598
9.6997 | -34.5626 | -19.571
9.7058 | -34.5632 | -19.582
9.7119 | -34.5638 | -19.593
9.718 | -34.5643 | -19.6044
9.7241 | -34.5649 | -19.6155
9.7302 | -34.5654 | -19.6267
9.7363 | -34.566 | -19.6376
9.7424 | -34.5666 | -19.6485
9.7485 | -34.5672 | -19.6594
9.7546 | -34.5678 | -19.6702
9.7607 | -34.5684 | -19.6815
9.7668 | -34.5689 | -19.6929
9.7729 | -34.5695 | -19.7037
9.779 | -34.5702 | -19.7147
9.7851 | -34.5707 | -19.7259
9.7912 | -34.5713 | -19.7375
9.7973 | -34.5718 | -19.7487
9.8034 | -34.5725 | -19.7599
9.8095 | -34.573 | -19.7709
9.8157 | -34.5736 | -19.7819
9.8218 | -34.5742 | -19.7927
9.8279 | -34.5748 | -19.8038
9.834 | -34.5754 | -19.8147
9.8401 | -34.576 | -19.8257
9.8462 | -34.5766 | -19.8365
9.8523 | -34.5772 | -19.8474
9.8584 | -34.5778 | -19.8587
9.8645 | -34.5783 | -19.8698
9.8706 | -34.5789 | -19.8807
9.8767 | -34.5794 | -19.8914
9.8828 | -34.5801 | -19.902
9.8889 | -34.5806 | -19.9132
9.895 | -34.5812 | -19.9239
9.9011 | -34.5818 | -19.9345
9.9072 | -34.5824 | -19.9453
9.9133 | -34.5829 | -19.9561
9.9194 | -34.5836 | -19.9668
9.9255 | -34.5841 | -19.9779
9.9316 | -34.5847 | -19.9887
9.9377 | -34.5854 | -19.9995
9.9438 | -34.5859 | -20.0104
9.9499 | -34.5865 | -20.0215
9.956 | -34.5871 | -20.0323
9.9622 | -34.5877 | -20.0431
9.9683 | -34.5883 | -20.0541
9.9744 | -34.5888 | -20.065
9.9805 | -34.5894 | -20.076
9.9866 | -34.59 | -20.0866
9.9927 | -34.5906 | -20.0978
9.9988 | -34.5912 | -20.1086
10.0049 | -34.5918 | -20.1195
10.011 | -34.5924 | -20.1304
10.0171 | -34.593 | -20.1412
10.0232 | -34.5936 | -20.1519
10.0293 | -34.5942 | -20.1627
10.0354 | -34.5949 | -20.1733
10.0415 | -34.5955 | -20.1842
10.0476 | -34.596 | -20.1954
10.0537 | -34.5966 | -20.2066
10.0598 | -34.5972 | -20.2176
10.0659 | -34.5977 | -20.2285
10.072 | -34.5984 | -20.2393
10.0781 | -34.599 | -20.2502
10.0842 | -34.5996 | -20.2612
10.0903 | -34.6002 | -20.2723
10.0964 | -34.6008 | -20.2829
10.1026 | -34.6014 | -20.2938
10.1087 | -34.602 | -20.3047
10.1148 | -34.6026 | -20.3156
10.1209 | -34.6032 | -20.3264
10.127 | -34.6038 | -20.337
10.1331 | -34.6044 | -20.3478
10.1392 | -34.605 | -20.3585
10.1453 | -34.6056 | -20.3694
10.1514 | -34.6062 | -20.38
10.1575 | -34.6068 | -20.3908
10.1636 | -34.6074 | -20.4009
10.1697 | -34.608 | -20.4117
10.1758 | -34.6086 | -20.4224
10.1819 | -34.6092 | -20.4333
10.188 | -34.6098 | -20.444
10.1941 | -34.6104 | -20.4549
10.2002 | -34.6109 | -20.4651
10.2063 | -34.6115 | -20.4762
10.2124 | -34.6121 | -20.487
10.2185 | -34.6127 | -20.4977
10.2246 | -34.6133 | -20.5085
10.2307 | -34.6139 | -20.5195
10.2368 | -34.6145 | -20.5294
10.2429 | -34.6151 | -20.54
10.2491 | -34.6157 | -20.5506
10.2552 | -34.6163 | -20.5617
10.2613 | -34.617 | -20.5727
10.2674 | -34.6175 | -20.5835
10.2735 | -34.6181 | -20.5946
10.2796 | -34.6187 | -20.6054
10.2857 | -34.6192 | -20.6163
10.2918 | -34.6198 | -20.6271
10.2979 | -34.6205 | -20.638
10.304 | -34.6211 | -20.649
10.3101 | -34.6217 | -20.6598
10.3162 | -34.6223 | -20.6705
10.3223 | -34.6229 | -20.6813
10.3284 | -34.6235 | -20.6919
10.3345 | -34.6241 | -20.7025
10.3406 | -34.6247 | -20.7131
10.3467 | -34.6253 | -20.7238
10.3528 | -34.6258 | -20.7343
10.3589 | -34.6264 | -20.7452
10.365 | -34.627 | -20.7561
10.3711 | -34.6276 | -20.7671
10.3772 | -34.6281 | -20.7779
10.3833 | -34.6287 | -20.7886
10.3895 | -34.6293 | -20.7992
10.3956 | -34.6299 | -20.8097
10.4017 | -34.6305 | -20.8203
10.4078 | -34.6311 | -20.8311
10.4139 | -34.6317 | -20.8418
10.42 | -34.6323 | -20.8526
10.4261 | -34.633 | -20.8632
10.4322 | -34.6336 | -20.8738
10.4383 | -34.6341 | -20.8845
10.4444 | -34.6347 | -20.8948
10.4505 | -34.6353 | -20.9056
10.4566 | -34.6359 | -20.9162
10.4627 | -34.6365 | -20.9269
10.4688 | -34.6371 | -20.9371
10.4749 | -34.6377 | -20.9477
10.481 | -34.6383 | -20.9583
10.4871 | -34.6389 | -20.9691
10.4932 | -34.6396 | -20.9797
10.4993 | -34.6402 | -20.9903
10.5054 | -34.6408 | -21.001
10.5115 | -34.6414 | -21.0117
10.5176 | -34.642 | -21.022
10.5237 | -34.6426 | -21.0326
10.5298 | -34.6432 | -21.043
10.536 | -34.6438 | -21.0538
10.5421 | -34.6444 | -21.0641
10.5482 | -34.645 | -21.075
10.5543 | -34.6457 | -21.0855
10.5604 | -34.6463 | -21.0958
10.5665 | -34.6469 | -21.1063
10.5726 | -34.6474 | -21.1167
10.5787 | -34.6481 | -21.1275
10.5848 | -34.6487 | -21.1377
10.5909 | -34.6493 | -21.1484
10.597 | -34.6498 | -21.159
10.6031 | -34.6505 | -21.1697
10.6092 | -34.6511 | -21.1805
10.6153 | -34.6516 | -21.1913
10.6214 | -34.6523 | -21.2017
10.6275 | -34.6529 | -21.2123
10.6336 | -34.6535 | -21.2232
10.6397 | -34.6541 | -21.2337
10.6458 | -34.6546 | -21.244
10.6519 | -34.6552 | -21.2546
10.658 | -34.6558 | -21.2649
10.6641 | -34.6564 | -21.2755
10.6702 | -34.657 | -21.2863
10.6764 | -34.6577 | -21.2968
10.6825 | -34.6583 | -21.3075
10.6886 | -34.6589 | -21.3181
10.6947 | -34.6595 | -21.3284
10.7008 | -34.6601 | -21.339
10.7069 | -34.6607 | -21.3496
10.713 | -34.6614 | -21.3601
10.7191 | -34.662 | -21.3708
10.7252 | -34.6626 | -21.3811
10.7313 | -34.6632 | -21.3916
10.7374 | -34.6638 | -21.402
10.7435 | -34.6644 | -21.4128
10.7496 | -34.665 | -21.4231
10.7557 | -34.6656 | -21.4337
10.7618 | -34.6662 | -21.4442
10.7679 | -34.6669 | -21.4547
10.774 | -34.6674 | -21.4652
10.7801 | -34.668 | -21.4761
10.7862 | -34.6686 | -21.4865
10.7923 | -34.6693 | -21.497
10.7984 | -34.6699 | -21.5074
10.8045 | -34.6705 | -21.518
10.8106 | -34.6711 | -21.5285
10.8168 | -34.6717 | -21.5391
10.8229 | -34.6723 | -21.5497
10.829 | -34.6729 | -21.5602
10.8351 | -34.6735 | -21.5707
10.8412 | -34.6741 | -21.5812
10.8473 | -34.6747 | -21.5914
10.8534 | -34.6753 | -21.602
10.8595 | -34.6759 | -21.6124
10.8656 | -34.6765 | -21.6226
10.8717 | -34.6772 | -21.6328
10.8778 | -34.6778 | -21.6431
10.8839 | -34.6784 | -21.6535
10.89 | -34.679 | -21.6638
10.8961 | -34.6796 | -21.6743
10.9022 | -34.6802 | -21.6846
10.9083 | -34.6808 | -21.6952
10.9144 | -34.6815 | -21.7056
10.9205 | -34.6821 | -21.716
10.9266 | -34.6827 | -21.7267
10.9327 | -34.6834 | -21.7372
10.9388 | -34.684 | -21.7476
10.9449 | -34.6846 | -21.7581
10.951 | -34.6852 | -21.7685
10.9571 | -34.6857 | -21.779
10.9633 | -34.6864 | -21.7897
10.9694 | -34.687 | -21.8005
10.9755 | -34.6876 | -21.8108
10.9816 | -34.6882 | -21.8213
10.9877 | -34.6889 | -21.8317
10.9938 | -34.6894 | -21.8424
10.9999 | -34.6901 | -21.8528
11.006 | -34.6907 | -21.863
11.0121 | -34.6913 | -21.8737
11.0182 | -34.6919 | -21.8839
11.0243 | -34.6925 | -21.8943
11.0304 | -34.6931 | -21.9046
11.0365 | -34.6937 | -21.915
11.0426 | -34.6943 | -21.9252
11.0487 | -34.695 | -21.9357
11.0548 | -34.6956 | -21.946
11.0609 | -34.6962 | -21.9567
11.067 | -34.6968 | -21.9673
11.0731 | -34.6975 | -21.9777
11.0792 | -34.6981 | -21.988
11.0853 | -34.6987 | -21.9982
11.0914 | -34.6993 | -22.0085
11.0975 | -34.6999 | -22.0191
11.1037 | -34.7005 | -22.0293
11.1098 | -34.7011 | -22.0399
11.1159 | -34.7018 | -22.0504
11.122 | -34.7024 | -22.061
11.1281 | -34.703 | -22.0712
11.1342 | -34.7036 | -22.0814
11.1403 | -34.7042 | -22.0921
11.1464 | -34.7049 | -22.1025
11.1525 | -34.7055 | -22.1129
11.1586 | -34.7061 | -22.1239
11.1647 | -34.7067 | -22.1341
11.1708 | -34.7073 | -22.1443
11.1769 | -34.7079 | -22.1543
11.183 | -34.7085 | -22.1646
11.1891 | -34.7092 | -22.1752
11.1952 | -34.7098 | -22.1859
11.2013 | -34.7103 | -22.1963
11.2074 | -34.7109 | -22.2064
11.2135 | -34.7115 | -22.2169
11.2196 | -34.7121 | -22.2271
11.2257 | -34.7127 | -22.2374
11.2318 | -34.7132 | -22.248
11.2379 | -34.7138 | -22.2583
11.244 | -34.7144 | -22.2686
11.2502 | -34.715 | -22.279
11.2563 | -34.7157 | -22.2894
11.2624 | -34.7163 | -22.2998
11.2685 | -34.7169 | -22.31
11.2746 | -34.7175 | -22.3201
11.2807 | -34.7181 | -22.3308
11.2868 | -34.7186 | -22.3415
11.2929 | -34.7193 | -22.3515
11.299 | -34.7199 | -22.3621
11.3051 | -34.7205 | -22.3725
11.3112 | -34.7211 | -22.3829
11.3173 | -34.7217 | -22.3933
11.3234 | -34.7223 | -22.4037
11.3295 | -34.7229 | -22.4138
11.3356 | -34.7236 | -22.4241
11.3417 | -34.7242 | -22.4346
11.3478 | -34.7248 | -22.4449
11.3539 | -34.7253 | -22.4552
11.36 | -34.7259 | -22.4654
11.3661 | -34.7266 | -22.4757
11.3722 | -34.7272 | -22.4859
11.3783 | -34.7278 | -22.4961
11.3844 | -34.7284 | -22.5065
11.3906 | -34.7289 | -22.5171
11.3967 | -34.7296 | -22.5271
11.4028 | -34.7301 | -22.5376
11.4089 | -34.7308 | -22.5478
11.415 | -34.7313 | -22.5581
11.4211 | -34.7319 | -22.5682
11.4272 | -34.7325 | -22.5786
11.4333 | -34.7331 | -22.5889
11.4394 | -34.7337 | -22.5995
11.4455 | -34.7343 | -22.6098
11.4516 | -34.7349 | -22.62
11.4577 | -34.7355 | -22.6306
11.4638 | -34.7362 | -22.6404
11.4699 | -34.7368 | -22.651
11.476 | -34.7373 | -22.6618
11.4821 | -34.738 | -22.6719
11.4882 | -34.7386 | -22.6824
11.4943 | -34.7392 | -22.6926
11.5004 | -34.7398 | -22.7029
11.5065 | -34.7405 | -22.7132
11.5126 | -34.7411 | -22.7232
11.5187 | -34.7417 | -22.7335
11.5248 | -34.7423 | -22.7438
11.5309 | -34.7429 | -22.7541
11.5371 | -34.7435 | -22.7643
11.5432 | -34.7441 | -22.7747
11.5493 | -34.7447 | -22.7852
11.5554 | -34.7453 | -22.7953
11.5615 | -34.7459 | -22.8058
11.5676 | -34.7465 | -22.8162
11.5737 | -34.7472 | -22.8262
11.5798 | -34.7478 | -22.8363
11.5859 | -34.7484 | -22.8464
11.592 | -34.749 | -22.8567
11.5981 | -34.7496 | -22.8668
11.6042 | -34.7503 | -22.8771
11.6103 | -34.7508 | -22.8873
11.6164 | -34.7515 | -22.8973
11.6225 | -34.7521 | -22.9076
11.6286 | -34.7527 | -22.9179
11.6347 | -34.7533 | -22.9284
11.6408 | -34.7539 | -22.9386
11.6469 | -34.7546 | -22.9491
11.653 | -34.7552 | -22.9591
11.6591 | -34.7558 | -22.9694
11.6652 | -34.7564 | -22.9798
11.6713 | -34.757 | -22.99
11.6775 | -34.7577 | -23
11.6836 | -34.7583 | -23.0104
11.6897 | -34.7589 | -23.0204
11.6958 | -34.7595 | -23.031
11.7019 | -34.7602 | -23.0413
11.708 | -34.7608 | -23.0518
11.7141 | -34.7614 | -23.062
11.7202 | -34.762 | -23.072
11.7263 | -34.7626 | -23.0823
11.7324 | -34.7633 | -23.0927
11.7385 | -34.7638 | -23.1031
11.7446 | -34.7644 | -23.1135
11.7507 | -34.7651 | -23.1237
11.7568 | -34.7657 | -23.134
11.7629 | -34.7663 | -23.1442
11.769 | -34.7669 | -23.1544
11.7751 | -34.7676 | -23.1644
11.7812 | -34.7682 | -23.1744
11.7873 | -34.7688 | -23.1848
11.7934 | -34.7695 | -23.1951
11.7995 | -34.7701 | -23.205
11.8056 | -34.7706 | -23.2156
11.8117 | -34.7712 | -23.2259
11.8178 | -34.7719 | -23.2361
11.824 | -34.7725 | -23.2461
11.8301 | -34.7732 | -23.256
11.8362 | -34.7738 | -23.2659
11.8423 | -34.7744 | -23.2761
11.8484 | -34.775 | -23.2863
11.8545 | -34.7756 | -23.2966
11.8606 | -34.7762 | -23.3068
11.8667 | -34.7769 | -23.3171
11.8728 | -34.7775 | -23.3273
11.8789 | -34.7781 | -23.3378
11.885 | -34.7786 | -23.3479
11.8911 | -34.7793 | -23.3583
11.8972 | -34.7799 | -23.3681
11.9033 | -34.7805 | -23.3779
11.9094 | -34.7811 | -23.3878
11.9155 | -34.7817 | -23.398
11.9216 | -34.7824 | -23.408
11.9277 | -34.783 | -23.418
11.9338 | -34.7836 | -23.4283
11.9399 | -34.7842 | -23.4388
11.946 | -34.7848 | -23.449
11.9521 | -34.7855 | -23.459
11.9582 | -34.7861 | -23.469
11.9644 | -34.7868 | -23.479
11.9705 | -34.7874 | -23.4894
11.9766 | -34.788 | -23.4995
11.9827 | -34.7886 | -23.5096
11.9888 | -34.7892 | -23.5196
11.9949 | -34.7898 | -23.5296
12.001 | -34.7905 | -23.5398
12.0071 | -34.7911 | -23.5497
12.0132 | -34.7917 | -23.5596
12.0193 | -34.7923 | -23.57
12.0254 | -34.793 | -23.5799
12.0315 | -34.7936 | -23.5899
12.0376 | -34.7942 | -23.5997
12.0437 | -34.7948 | -23.6097
12.0498 | -34.7954 | -23.62
12.0559 | -34.7961 | -23.6302
12.062 | -34.7967 | -23.6402
12.0681 | -34.7973 | -23.6502
12.0742 | -34.798 | -23.6605
12.0803 | -34.7986 | -23.6705
12.0864 | -34.7992 | -23.6806
12.0925 | -34.7999 | -23.6907
12.0986 | -34.8005 | -23.7005
12.1047 | -34.8012 | -23.7106
12.1109 | -34.8018 | -23.7208
12.117 | -34.8025 | -23.7308
12.1231 | -34.8031 | -23.7409
12.1292 | -34.8037 | -23.7511
12.1353 | -34.8044 | -23.7609
12.1414 | -34.805 | -23.7711
12.1475 | -34.8056 | -23.7812
12.1536 | -34.8062 | -23.7914
12.1597 | -34.8069 | -23.8015
12.1658 | -34.8075 | -23.8115
12.1719 | -34.8082 | -23.8216
12.178 | -34.8088 | -23.8318
12.1841 | -34.8094 | -23.842
12.1902 | -34.8101 | -23.8521
12.1963 | -34.8107 | -23.8621
12.2024 | -34.8114 | -23.872
12.2085 | -34.812 | -23.8821
12.2146 | -34.8126 | -23.892
12.2207 | -34.8133 | -23.9021
12.2268 | -34.8139 | -23.9122
12.2329 | -34.8145 | -23.9222
12.239 | -34.8151 | -23.9319
12.2451 | -34.8158 | -23.9418
12.2513 | -34.8164 | -23.9519
12.2574 | -34.8171 | -23.9622
12.2635 | -34.8177 | -23.9718
12.2696 | -34.8183 | -23.9817
12.2757 | -34.819 | -23.9913
12.2818 | -34.8196 | -24.0012
12.2879 | -34.8202 | -24.0111
12.294 | -34.8208 | -24.0209
12.3001 | -34.8215 | -24.0311
12.3062 | -34.8221 | -24.0411
12.3123 | -34.8228 | -24.051
12.3184 | -34.8234 | -24.0608
12.3245 | -34.8241 | -24.0709
12.3306 | -34.8247 | -24.0807
12.3367 | -34.8253 | -24.0905
12.3428 | -34.826 | -24.1003
12.3489 | -34.8266 | -24.1097
12.355 | -34.8273 | -24.1196
12.3611 | -34.8279 | -24.1296
12.3672 | -34.8285 | -24.1393
12.3733 | -34.8291 | -24.1496
12.3794 | -34.8298 | -24.1591
12.3855 | -34.8304 | -24.1688
12.3916 | -34.831 | -24.1789
12.3978 | -34.8317 | -24.1891
12.4039 | -34.8323 | -24.199
12.41 | -34.8329 | -24.2091
12.4161 | -34.8336 | -24.2187
12.4222 | -34.8342 | -24.2285
12.4283 | -34.8349 | -24.2384
12.4344 | -34.8355 | -24.2484
12.4405 | -34.8362 | -24.2584
12.4466 | -34.8368 | -24.2682
12.4527 | -34.8374 | -24.278
12.4588 | -34.8381 | -24.2879
12.4649 | -34.8387 | -24.2981
12.471 | -34.8394 | -24.3084
12.4771 | -34.84 | -24.3179
12.4832 | -34.8407 | -24.3277
12.4893 | -34.8413 | -24.3374
12.4954 | -34.842 | -24.3472
12.5015 | -34.8427 | -24.3569
12.5076 | -34.8433 | -24.3668
12.5137 | -34.8439 | -24.377
12.5198 | -34.8445 | -24.3866
12.5259 | -34.8452 | -24.3963
12.532 | -34.8459 | -24.4063
12.5382 | -34.8465 | -24.4161
12.5443 | -34.8471 | -24.426
12.5504 | -34.8478 | -24.436
12.5565 | -34.8485 | -24.4456
12.5626 | -34.8491 | -24.4556
12.5687 | -34.8498 | -24.4656
12.5748 | -34.8504 | -24.4754
12.5809 | -34.851 | -24.4849
12.587 | -34.8516 | -24.4951
12.5931 | -34.8523 | -24.5049
12.5992 | -34.8529 | -24.5149
12.6053 | -34.8536 | -24.5246
12.6114 | -34.8542 | -24.5345
12.6175 | -34.8549 | -24.5444
12.6236 | -34.8555 | -24.554
12.6297 | -34.8562 | -24.5636
12.6358 | -34.8569 | -24.5732
12.6419 | -34.8576 | -24.5833
12.648 | -34.8582 | -24.5933
12.6541 | -34.8588 | -24.6031
12.6602 | -34.8595 | -24.6124
12.6663 | -34.8601 | -24.6221
12.6724 | -34.8608 | -24.6319
12.6785 | -34.8614 | -24.6419
12.6847 | -34.8621 | -24.6518
12.6908 | -34.8627 | -24.6614
12.6969 | -34.8634 | -24.6711
12.703 | -34.8641 | -24.681
12.7091 | -34.8647 | -24.6908
12.7152 | -34.8654 | -24.7005
12.7213 | -34.866 | -24.7104
12.7274 | -34.8666 | -24.7203
12.7335 | -34.8673 | -24.7294
12.7396 | -34.8679 | -24.7388
12.7457 | -34.8686 | -24.7486
12.7518 | -34.8692 | -24.7584
12.7579 | -34.8698 | -24.7685
12.764 | -34.8705 | -24.7782
12.7701 | -34.8711 | -24.7876
12.7762 | -34.8717 | -24.7972
12.7823 | -34.8724 | -24.807
12.7884 | -34.873 | -24.8168
12.7945 | -34.8736 | -24.8262
12.8006 | -34.8743 | -24.8352
12.8067 | -34.8749 | -24.8451
12.8128 | -34.8755 | -24.8545
12.8189 | -34.8761 | -24.8639
12.8251 | -34.8767 | -24.8734
12.8312 | -34.8773 | -24.8832
12.8373 | -34.878 | -24.8929
12.8434 | -34.8787 | -24.9025
12.8495 | -34.8793 | -24.9121
12.8556 | -34.8799 | -24.9217
12.8617 | -34.8805 | -24.9312
12.8678 | -34.8812 | -24.9409
12.8739 | -34.8818 | -24.9502
12.88 | -34.8824 | -24.9599
12.8861 | -34.883 | -24.9697
12.8922 | -34.8837 | -24.9793
12.8983 | -34.8844 | -24.9887
12.9044 | -34.885 | -24.9982
12.9105 | -34.8857 | -25.0076
12.9166 | -34.8864 | -25.0176
12.9227 | -34.887 | -25.0277
12.9288 | -34.8877 | -25.0375
12.9349 | -34.8884 | -25.0474
12.941 | -34.8891 | -25.057
12.9471 | -34.8897 | -25.0667
12.9532 | -34.8904 | -25.0763
12.9593 | -34.8911 | -25.0862
12.9654 | -34.8917 | -25.096
12.9716 | -34.8923 | -25.1057
12.9777 | -34.8929 | -25.1149
12.9838 | -34.8936 | -25.1243
12.9899 | -34.8942 | -25.134
12.996 | -34.8948 | -25.1438
13.0021 | -34.8954 | -25.1535
13.0082 | -34.8961 | -25.163
13.0143 | -34.8967 | -25.1723
13.0204 | -34.8973 | -25.1816
13.0265 | -34.898 | -25.1909
13.0326 | -34.8986 | -25.2007
13.0387 | -34.8993 | -25.2105
13.0448 | -34.8999 | -25.2201
13.0509 | -34.9005 | -25.2297
13.057 | -34.9012 | -25.2389
13.0631 | -34.9019 | -25.2483
13.0692 | -34.9026 | -25.2578
13.0753 | -34.9033 | -25.2673
13.0814 | -34.9039 | -25.2767
13.0875 | -34.9045 | -25.2858
13.0936 | -34.9052 | -25.2952
13.0997 | -34.9058 | -25.3044
13.1058 | -34.9064 | -25.3134
13.112 | -34.907 | -25.323
13.1181 | -34.9077 | -25.3324
13.1242 | -34.9083 | -25.342
13.1303 | -34.909 | -25.3517
13.1364 | -34.9096 | -25.3612
13.1425 | -34.9103 | -25.3706
13.1486 | -34.9109 | -25.3801
13.1547 | -34.9116 | -25.3891
13.1608 | -34.9122 | -25.3985
13.1669 | -34.9129 | -25.4079
13.173 | -34.9135 | -25.4173
13.1791 | -34.9141 | -25.4268
13.1852 | -34.9148 | -25.4365
13.1913 | -34.9154 | -25.4459
13.1974 | -34.916 | -25.4554
13.2035 | -34.9167 | -25.4648
13.2096 | -34.9174 | -25.4742
13.2157 | -34.918 | -25.4837
13.2218 | -34.9187 | -25.4933
13.2279 | -34.9194 | -25.5025
13.234 | -34.92 | -25.512
13.2401 | -34.9207 | -25.5214
13.2462 | -34.9213 | -25.5308
13.2524 | -34.9219 | -25.5401
13.2585 | -34.9226 | -25.5494
13.2646 | -34.9232 | -25.559
13.2707 | -34.9239 | -25.5686
13.2768 | -34.9245 | -25.5784
13.2829 | -34.9251 | -25.5879
13.289 | -34.9258 | -25.5973
13.2951 | -34.9264 | -25.6065
13.3012 | -34.9271 | -25.616
13.3073 | -34.9277 | -25.6257
13.3134 | -34.9283 | -25.635
13.3195 | -34.9289 | -25.6443
13.3256 | -34.9296 | -25.6537
13.3317 | -34.9302 | -25.6629
13.3378 | -34.9309 | -25.6725
13.3439 | -34.9315 | -25.6819
13.35 | -34.9322 | -25.6913
13.3561 | -34.9328 | -25.7009
13.3622 | -34.9334 | -25.7103
13.3683 | -34.9341 | -25.7194
13.3744 | -34.9347 | -25.7287
13.3805 | -34.9353 | -25.7383
13.3866 | -34.9359 | -25.7474
13.3927 | -34.9365 | -25.7568
13.3989 | -34.9372 | -25.7657
13.405 | -34.9379 | -25.775
13.4111 | -34.9385 | -25.7844
13.4172 | -34.9392 | -25.7935
13.4233 | -34.9398 | -25.8031
13.4294 | -34.9404 | -25.8123
13.4355 | -34.9411 | -25.8221
13.4416 | -34.9417 | -25.8314
13.4477 | -34.9423 | -25.8408
13.4538 | -34.9429 | -25.8502
13.4599 | -34.9436 | -25.8597
13.466 | -34.9442 | -25.8694
13.4721 | -34.9449 | -25.8787
13.4782 | -34.9455 | -25.888
13.4843 | -34.9462 | -25.8974
13.4904 | -34.9468 | -25.9062
13.4965 | -34.9474 | -25.9159
13.5026 | -34.9481 | -25.9253
13.5087 | -34.9487 | -25.9342
13.5148 | -34.9493 | -25.9437
13.5209 | -34.95 | -25.9529
13.527 | -34.9506 | -25.9625
13.5331 | -34.9512 | -25.9717
13.5393 | -34.9519 | -25.9811
13.5454 | -34.9525 | -25.9904
13.5515 | -34.9532 | -25.9998
13.5576 | -34.9538 | -26.0092
13.5637 | -34.9545 | -26.0185
13.5698 | -34.9551 | -26.028
13.5759 | -34.9557 | -26.037
13.582 | -34.9563 | -26.0465
13.5881 | -34.9569 | -26.0557
13.5942 | -34.9575 | -26.0651
13.6003 | -34.9582 | -26.0745
13.6064 | -34.9588 | -26.084
13.6125 | -34.9595 | -26.0934
13.6186 | -34.9602 | -26.1027
13.6247 | -34.9608 | -26.112
13.6308 | -34.9614 | -26.1211
13.6369 | -34.962 | -26.1304
13.643 | -34.9627 | -26.1395
13.6491 | -34.9633 | -26.1488
13.6552 | -34.964 | -26.158
13.6613 | -34.9646 | -26.1675
13.6674 | -34.9652 | -26.1769
13.6735 | -34.9658 | -26.1863
13.6796 | -34.9665 | -26.1954
13.6858 | -34.967 | -26.2047
13.6919 | -34.9677 | -26.2142
13.698 | -34.9683 | -26.2236
13.7041 | -34.9689 | -26.2332
13.7102 | -34.9696 | -26.2427
13.7163 | -34.9703 | -26.252
13.7224 | -34.9709 | -26.2612
13.7285 | -34.9716 | -26.2706
13.7346 | -34.9722 | -26.2799
13.7407 | -34.9728 | -26.2893
13.7468 | -34.9735 | -26.2986
13.7529 | -34.9741 | -26.3076
13.759 | -34.9747 | -26.3171
13.7651 | -34.9753 | -26.3265
13.7712 | -34.9759 | -26.3359
13.7773 | -34.9766 | -26.3454
13.7834 | -34.9772 | -26.3547
13.7895 | -34.9779 | -26.3639
13.7956 | -34.9785 | -26.3732
13.8017 | -34.9792 | -26.3824
13.8078 | -34.9798 | -26.3916
13.8139 | -34.9805 | -26.4007
13.82 | -34.9811 | -26.4097
13.8262 | -34.9817 | -26.4193
13.8323 | -34.9824 | -26.4288
13.8384 | -34.983 | -26.4383
13.8445 | -34.9837 | -26.4477
13.8506 | -34.9843 | -26.457
13.8567 | -34.9849 | -26.4665
13.8628 | -34.9855 | -26.4758
13.8689 | -34.9861 | -26.4852
13.875 | -34.9868 | -26.4945
13.8811 | -34.9875 | -26.5038
13.8872 | -34.9881 | -26.5131
13.8933 | -34.9888 | -26.5222
13.8994 | -34.9894 | -26.5312
13.9055 | -34.99 | -26.5405
13.9116 | -34.9907 | -26.5496
13.9177 | -34.9913 | -26.5586
13.9238 | -34.992 | -26.568
13.9299 | -34.9927 | -26.5771
13.936 | -34.9932 | -26.5861
13.9421 | -34.9939 | -26.5953
13.9482 | -34.9945 | -26.6046
13.9543 | -34.9952 | -26.6137
13.9604 | -34.9958 | -26.6231
13.9665 | -34.9965 | -26.6324
13.9727 | -34.9971 | -26.6415
13.9788 | -34.9977 | -26.651
13.9849 | -34.9983 | -26.6603
13.991 | -34.9989 | -26.6697
13.9971 | -34.9995 | -26.6791
14.0032 | -35.0001 | -26.6885
14.0093 | -35.0007 | -26.6976
14.0154 | -35.0013 | -26.7072
14.0215 | -35.0019 | -26.7162
14.0276 | -35.0026 | -26.7252
14.0337 | -35.0033 | -26.7346
14.0398 | -35.0039 | -26.7437
14.0459 | -35.0045 | -26.7531
14.052 | -35.0051 | -26.7624
14.0581 | -35.0057 | -26.7714
14.0642 | -35.0064 | -26.7809
14.0703 | -35.007 | -26.7899
14.0764 | -35.0076 | -26.7995
14.0825 | -35.0083 | -26.8088
14.0886 | -35.0089 | -26.8177
14.0947 | -35.0095 | -26.8267
14.1008 | -35.0101 | -26.8363
14.1069 | -35.0107 | -26.8454
14.1131 | -35.0113 | -26.8549
14.1192 | -35.0119 | -26.8638
14.1253 | -35.0125 | -26.8728
14.1314 | -35.0132 | -26.8819
14.1375 | -35.0139 | -26.8912
14.1436 | -35.0146 | -26.9004
14.1497 | -35.0152 | -26.9094
14.1558 | -35.0158 | -26.9188
14.1619 | -35.0165 | -26.928
14.168 | -35.0171 | -26.9371
14.1741 | -35.0178 | -26.9461
14.1802 | -35.0184 | -26.9552
14.1863 | -35.019 | -26.9645
14.1924 | -35.0197 | -26.9738
14.1985 | -35.0203 | -26.9828
14.2046 | -35.0209 | -26.9918
14.2107 | -35.0216 | -27.0006
14.2168 | -35.0222 | -27.0102
14.2229 | -35.0229 | -27.0195
14.229 | -35.0235 | -27.0287
14.2351 | -35.0241 | -27.0377
14.2412 | -35.0248 | -27.0466
14.2473 | -35.0255 | -27.0559
14.2534 | -35.0261 | -27.0649
14.2596 | -35.0267 | -27.0742
14.2657 | -35.0273 | -27.0834
14.2718 | -35.028 | -27.0927
14.2779 | -35.0286 | -27.1017
14.284 | -35.0292 | -27.1108
14.2901 | -35.0299 | -27.1198
14.2962 | -35.0306 | -27.1288
14.3023 | -35.0312 | -27.1381
14.3084 | -35.0318 | -27.1475
14.3145 | -35.0325 | -27.1566
14.3206 | -35.0331 | -27.1658
14.3267 | -35.0337 | -27.1749
14.3328 | -35.0343 | -27.1841
14.3389 | -35.0349 | -27.193
14.345 | -35.0356 | -27.2023
14.3511 | -35.0362 | -27.2112
14.3572 | -35.0368 | -27.2205
14.3633 | -35.0375 | -27.2296
14.3694 | -35.0381 | -27.2388
14.3755 | -35.0387 | -27.2477
14.3816 | -35.0393 | -27.2568
14.3877 | -35.04 | -27.2657
14.3938 | -35.0406 | -27.2752
14.4 | -35.0413 | -27.284
14.4061 | -35.0419 | -27.2931
14.4122 | -35.0425 | -27.3024
14.4183 | -35.0431 | -27.3114
14.4244 | -35.0437 | -27.3203
14.4305 | -35.0443 | -27.3296
14.4366 | -35.045 | -27.3389
14.4427 | -35.0456 | -27.3483
14.4488 | -35.0462 | -27.3572
14.4549 | -35.0469 | -27.3662
14.461 | -35.0475 | -27.3755
14.4671 | -35.0482 | -27.3846
14.4732 | -35.0488 | -27.3934
14.4793 | -35.0495 | -27.4032
14.4854 | -35.0502 | -27.412
14.4915 | -35.0508 | -27.4209
14.4976 | -35.0515 | -27.4301
14.5037 | -35.0521 | -27.4393
14.5098 | -35.0527 | -27.4483
14.5159 | -35.0534 | -27.4575
14.522 | -35.054 | -27.4667
14.5281 | -35.0546 | -27.4754
14.5342 | -35.0553 | -27.4847
14.5403 | -35.0559 | -27.4938
14.5465 | -35.0566 | -27.5032
14.5526 | -35.0572 | -27.5126
14.5587 | -35.0578 | -27.5216
14.5648 | -35.0585 | -27.5304
14.5709 | -35.0591 | -27.5396
14.577 | -35.0597 | -27.5487
14.5831 | -35.0604 | -27.5579
14.5892 | -35.0611 | -27.5667
14.5953 | -35.0617 | -27.5754
14.6014 | -35.0623 | -27.5845
14.6075 | -35.063 | -27.5932
14.6136 | -35.0636 | -27.6021
14.6197 | -35.0642 | -27.6113
14.6258 | -35.0649 | -27.6205
14.6319 | -35.0656 | -27.6293
14.638 | -35.0662 | -27.6383
14.6441 | -35.0668 | -27.6469
14.6502 | -35.0675 | -27.6557
14.6563 | -35.0681 | -27.6646
14.6624 | -35.0688 | -27.6736
14.6685 | -35.0694 | -27.6822
14.6746 | -35.07 | -27.6914
14.6807 | -35.0707 | -27.7004
14.6869 | -35.0713 | -27.7093
14.693 | -35.0719 | -27.7183
14.6991 | -35.0725 | -27.7271
14.7052 | -35.0732 | -27.736
14.7113 | -35.0738 | -27.7453
14.7174 | -35.0745 | -27.7543
14.7235 | -35.0751 | -27.7632
14.7296 | -35.0757 | -27.7722
14.7357 | -35.0764 | -27.781
14.7418 | -35.077 | -27.7897
14.7479 | -35.0777 | -27.7987
14.754 | -35.0784 | -27.8076
14.7601 | -35.079 | -27.8166
14.7662 | -35.0796 | -27.8255
14.7723 | -35.0803 | -27.8342
14.7784 | -35.0809 | -27.8434
14.7845 | -35.0815 | -27.8525
14.7906 | -35.0822 | -27.8614
14.7967 | -35.0828 | -27.8702
14.8028 | -35.0834 | -27.8789
14.8089 | -35.0841 | -27.8878
14.815 | -35.0847 | -27.8965
14.8211 | -35.0853 | -27.9055
14.8272 | -35.0859 | -27.9147
14.8334 | -35.0866 | -27.9235
14.8395 | -35.0872 | -27.9324
14.8456 | -35.0879 | -27.9412
14.8517 | -35.0885 | -27.9503
14.8578 | -35.0892 | -27.9593
14.8639 | -35.0898 | -27.9684
14.87 | -35.0905 | -27.9774
14.8761 | -35.0912 | -27.986
14.8822 | -35.0917 | -27.995
14.8883 | -35.0924 | -28.0038
14.8944 | -35.0931 | -28.0129
14.9005 | -35.0937 | -28.0217
14.9066 | -35.0944 | -28.0306
14.9127 | -35.095 | -28.0393
14.9188 | -35.0957 | -28.0485
14.9249 | -35.0963 | -28.0574
14.931 | -35.097 | -28.0661
14.9371 | -35.0976 | -28.0749
14.9432 | -35.0983 | -28.084
14.9493 | -35.0989 | -28.0926
14.9554 | -35.0995 | -28.1015
14.9615 | -35.1002 | -28.1105
14.9676 | -35.1008 | -28.1197
14.9738 | -35.1015 | -28.1284
14.9799 | -35.1021 | -28.1378
14.986 | -35.1027 | -28.1463
14.9921 | -35.1034 | -28.155
14.9982 | -35.1041 | -28.1642
15.0043 | -35.1047 | -28.1731
15.0104 | -35.1053 | -28.1819
15.0165 | -35.106 | -28.1909
15.0226 | -35.1066 | -28.1994
15.0287 | -35.1073 | -28.2087
15.0348 | -35.1079 | -28.2177
15.0409 | -35.1086 | -28.2268
15.047 | -35.1092 | -28.2358
15.0531 | -35.11 | -28.2448
15.0592 | -35.1106 | -28.2539
15.0653 | -35.1113 | -28.2628
15.0714 | -35.1119 | -28.2714
15.0775 | -35.1126 | -28.28
15.0836 | -35.1133 | -28.2888
15.0897 | -35.114 | -28.2975
15.0958 | -35.1146 | -28.3062
15.1019 | -35.1152 | -28.3151
15.108 | -35.1158 | -28.324
15.1141 | -35.1165 | -28.3328
15.1203 | -35.1172 | -28.3417
15.1264 | -35.1179 | -28.3506
15.1325 | -35.1185 | -28.3594
15.1386 | -35.1192 | -28.3683
15.1447 | -35.1198 | -28.3771
15.1508 | -35.1205 | -28.3863
15.1569 | -35.1212 | -28.3953
15.163 | -35.1219 | -28.4038
15.1691 | -35.1226 | -28.4126
15.1752 | -35.1232 | -28.4214
15.1813 | -35.1239 | -28.4303
15.1874 | -35.1246 | -28.4392
15.1935 | -35.1253 | -28.4477
15.1996 | -35.126 | -28.4567
15.2057 | -35.1266 | -28.4654
15.2118 | -35.1272 | -28.4741
15.2179 | -35.1279 | -28.4828
15.224 | -35.1286 | -28.4917
15.2301 | -35.1293 | -28.5002
15.2362 | -35.1299 | -28.509
15.2423 | -35.1306 | -28.518
15.2484 | -35.1312 | -28.5267
15.2545 | -35.1319 | -28.5355
15.2607 | -35.1326 | -28.5443
15.2668 | -35.1333 | -28.5528
15.2729 | -35.134 | -28.5614
15.279 | -35.1347 | -28.5704
15.2851 | -35.1353 | -28.5792
15.2912 | -35.136 | -28.5879
15.2973 | -35.1366 | -28.5966
15.3034 | -35.1373 | -28.6052
15.3095 | -35.1379 | -28.6138
15.3156 | -35.1386 | -28.6226
15.3217 | -35.1392 | -28.6312
15.3278 | -35.1399 | -28.6402
15.3339 | -35.1406 | -28.6488
15.34 | -35.1413 | -28.6574
15.3461 | -35.1419 | -28.6663
15.3522 | -35.1425 | -28.675
15.3583 | -35.1432 | -28.6838
15.3644 | -35.1439 | -28.6926
15.3705 | -35.1446 | -28.7012
15.3766 | -35.1452 | -28.7102
15.3827 | -35.1459 | -28.7188
15.3888 | -35.1466 | -28.7271
15.3949 | -35.1472 | -28.7356
15.401 | -35.1479 | -28.7445
15.4072 | -35.1486 | -28.7529
15.4133 | -35.1492 | -28.7617
15.4194 | -35.1499 | -28.7704
15.4255 | -35.1506 | -28.7792
15.4316 | -35.1513 | -28.788
15.4377 | -35.1519 | -28.7966
15.4438 | -35.1527 | -28.8049
15.4499 | -35.1533 | -28.8136
15.456 | -35.1539 | -28.8223
15.4621 | -35.1546 | -28.8311
15.4682 | -35.1553 | -28.84
15.4743 | -35.1561 | -28.8486
15.4804 | -35.1567 | -28.8573
15.4865 | -35.1574 | -28.866
15.4926 | -35.158 | -28.8749
15.4987 | -35.1587 | -28.8838
15.5048 | -35.1594 | -28.8925
15.5109 | -35.1601 | -28.9008
15.517 | -35.1608 | -28.9096
15.5231 | -35.1615 | -28.9183
15.5292 | -35.1621 | -28.9268
15.5353 | -35.1628 | -28.9356
15.5414 | -35.1635 | -28.9445
15.5476 | -35.1642 | -28.9533
15.5537 | -35.1648 | -28.9619
15.5598 | -35.1655 | -28.9703
15.5659 | -35.1662 | -28.9792
15.572 | -35.1669 | -28.9876
15.5781 | -35.1676 | -28.9964
15.5842 | -35.1683 | -29.0047
15.5903 | -35.1689 | -29.013
15.5964 | -35.1696 | -29.0215
15.6025 | -35.1703 | -29.0301
15.6086 | -35.171 | -29.0388
15.6147 | -35.1716 | -29.0474
15.6208 | -35.1724 | -29.0559
15.6269 | -35.1731 | -29.0645
15.633 | -35.1737 | -29.0731
15.6391 | -35.1744 | -29.0819
15.6452 | -35.1751 | -29.0905
15.6513 | -35.1758 | -29.0993
15.6574 | -35.1765 | -29.1078
15.6635 | -35.1771 | -29.1163
15.6696 | -35.1778 | -29.1249
15.6757 | -35.1785 | -29.1339
15.6818 | -35.1792 | -29.1427
15.688 | -35.1799 | -29.1512
15.6941 | -35.1806 | -29.1598
15.7002 | -35.1813 | -29.1685
15.7063 | -35.182 | -29.1772
15.7124 | -35.1826 | -29.1859
15.7185 | -35.1833 | -29.1946
15.7246 | -35.184 | -29.2031
15.7307 | -35.1847 | -29.2119
15.7368 | -35.1854 | -29.221
15.7429 | -35.1861 | -29.2294
15.749 | -35.1868 | -29.238
15.7551 | -35.1874 | -29.2465
15.7612 | -35.1881 | -29.2545
15.7673 | -35.1888 | -29.263
15.7734 | -35.1894 | -29.2715
15.7795 | -35.1901 | -29.2804
15.7856 | -35.1908 | -29.2889
15.7917 | -35.1915 | -29.2973
15.7978 | -35.1922 | -29.3059
15.8039 | -35.1928 | -29.3143
15.81 | -35.1935 | -29.3227
15.8161 | -35.1942 | -29.3314
15.8222 | -35.1949 | -29.34
15.8283 | -35.1956 | -29.3483
15.8345 | -35.1963 | -29.3569
15.8406 | -35.197 | -29.3654
15.8467 | -35.1977 | -29.374
15.8528 | -35.1984 | -29.3824
15.8589 | -35.1991 | -29.3909
15.865 | -35.1998 | -29.3993
15.8711 | -35.2005 | -29.4081
15.8772 | -35.2012 | -29.4167
15.8833 | -35.2019 | -29.4256
15.8894 | -35.2026 | -29.4341
15.8955 | -35.2033 | -29.4427
15.9016 | -35.204 | -29.4512
15.9077 | -35.2047 | -29.4597
15.9138 | -35.2054 | -29.4681
15.9199 | -35.2061 | -29.4767
15.926 | -35.2068 | -29.4855
15.9321 | -35.2075 | -29.4941
15.9382 | -35.2082 | -29.5025
15.9443 | -35.2089 | -29.5113
15.9504 | -35.2096 | -29.5198
15.9565 | -35.2103 | -29.5284
15.9626 | -35.211 | -29.537
15.9687 | -35.2117 | -29.5453
15.9749 | -35.2124 | -29.5539
15.981 | -35.2131 | -29.5626
15.9871 | -35.2138 | -29.5711
15.9932 | -35.2145 | -29.5798
15.9993 | -35.2152 | -29.5882
16.0054 | -35.2159 | -29.5967
16.0115 | -35.2166 | -29.6051
16.0176 | -35.2173 | -29.6138
16.0237 | -35.218 | -29.6227
16.0298 | -35.2187 | -29.6307
16.0359 | -35.2194 | -29.6392
16.042 | -35.2201 | -29.6476
16.0481 | -35.2208 | -29.6562
16.0542 | -35.2215 | -29.6649
16.0603 | -35.2222 | -29.6731
16.0664 | -35.2229 | -29.6816
16.0725 | -35.2236 | -29.6896
16.0786 | -35.2243 | -29.6983
16.0847 | -35.225 | -29.7067
16.0908 | -35.2257 | -29.7153
16.0969 | -35.2264 | -29.7237
16.103 | -35.227 | -29.7325
16.1091 | -35.2277 | -29.7408
16.1152 | -35.2284 | -29.7494
16.1214 | -35.2292 | -29.7579
16.1275 | -35.2298 | -29.7663
16.1336 | -35.2306 | -29.7747
16.1397 | -35.2313 | -29.7833
16.1458 | -35.232 | -29.7917
16.1519 | -35.2327 | -29.8004
16.158 | -35.2333 | -29.809
16.1641 | -35.234 | -29.8175
16.1702 | -35.2347 | -29.8262
16.1763 | -35.2354 | -29.8347
16.1824 | -35.2362 | -29.8429
16.1885 | -35.2369 | -29.8514
16.1946 | -35.2376 | -29.8599
16.2007 | -35.2383 | -29.8689
16.2068 | -35.239 | -29.8773
16.2129 | -35.2397 | -29.8858
16.219 | -35.2403 | -29.8941
16.2251 | -35.241 | -29.9027
16.2312 | -35.2417 | -29.9113
16.2373 | -35.2424 | -29.9197
16.2434 | -35.2431 | -29.9282
16.2495 | -35.2438 | -29.9364
16.2556 | -35.2446 | -29.9451
16.2618 | -35.2453 | -29.9533
16.2679 | -35.246 | -29.9617
16.274 | -35.2467 | -29.9703
16.2801 | -35.2474 | -29.9787
16.2862 | -35.2481 | -29.9871
16.2923 | -35.2488 | -29.9958
16.2984 | -35.2495 | -30.0044
16.3045 | -35.2502 | -30.0127
16.3106 | -35.2509 | -30.0208
16.3167 | -35.2516 | -30.0294
16.3228 | -35.2523 | -30.0378
16.3289 | -35.253 | -30.0463
16.335 | -35.2537 | -30.0549
16.3411 | -35.2543 | -30.0633
16.3472 | -35.255 | -30.0718
16.3533 | -35.2557 | -30.0803
16.3594 | -35.2564 | -30.0885
16.3655 | -35.2571 | -30.097
16.3716 | -35.2578 | -30.1053
16.3777 | -35.2585 | -30.1138
16.3838 | -35.2592 | -30.1222
16.3899 | -35.2599 | -30.1308
16.396 | -35.2605 | -30.1395
16.4021 | -35.2613 | -30.1479
16.4083 | -35.262 | -30.1563
16.4144 | -35.2627 | -30.1647
16.4205 | -35.2634 | -30.1731
16.4266 | -35.2641 | -30.1816
16.4327 | -35.2648 | -30.1899
16.4388 | -35.2655 | -30.1983
16.4449 | -35.2662 | -30.2067
16.451 | -35.2668 | -30.2153
16.4571 | -35.2675 | -30.224
16.4632 | -35.2682 | -30.2321
16.4693 | -35.2689 | -30.2405
16.4754 | -35.2696 | -30.2489
16.4815 | -35.2703 | -30.2575
16.4876 | -35.2709 | -30.2658
16.4937 | -35.2716 | -30.2742
16.4998 | -35.2723 | -30.2827
16.5059 | -35.273 | -30.2911
16.512 | -35.2737 | -30.2997
16.5181 | -35.2745 | -30.3082
16.5242 | -35.2752 | -30.3168
16.5303 | -35.2758 | -30.3252
16.5364 | -35.2765 | -30.3334
16.5425 | -35.2772 | -30.3419
16.5487 | -35.2779 | -30.3502
16.5548 | -35.2786 | -30.3585
16.5609 | -35.2793 | -30.3669
16.567 | -35.28 | -30.3753
16.5731 | -35.2807 | -30.3837
16.5792 | -35.2814 | -30.392
16.5853 | -35.2821 | -30.4002
16.5914 | -35.2828 | -30.4087
16.5975 | -35.2835 | -30.4173
16.6036 | -35.2842 | -30.4256
16.6097 | -35.2849 | -30.4339
16.6158 | -35.2856 | -30.4425
16.6219 | -35.2863 | -30.4506
16.628 | -35.287 | -30.4592
16.6341 | -35.2877 | -30.4674
16.6402 | -35.2884 | -30.4752
16.6463 | -35.2891 | -30.4836
16.6524 | -35.2897 | -30.4919
16.6585 | -35.2904 | -30.5004
16.6646 | -35.2912 | -30.5087
16.6707 | -35.2919 | -30.5168
16.6768 | -35.2925 | -30.525
16.6829 | -35.2932 | -30.5334
16.689 | -35.2939 | -30.5419
16.6952 | -35.2946 | -30.5502
16.7013 | -35.2953 | -30.5585
16.7074 | -35.296 | -30.5668
16.7135 | -35.2966 | -30.575
16.7196 | -35.2973 | -30.5834
16.7257 | -35.298 | -30.5917
16.7318 | -35.2987 | -30.6001
16.7379 | -35.2994 | -30.6085
16.744 | -35.3001 | -30.6162
16.7501 | -35.3008 | -30.6247
16.7562 | -35.3015 | -30.6331
16.7623 | -35.3022 | -30.6415
16.7684 | -35.3029 | -30.65
16.7745 | -35.3036 | -30.658
16.7806 | -35.3043 | -30.6662
16.7867 | -35.305 | -30.6745
16.7928 | -35.3057 | -30.6827
16.7989 | -35.3063 | -30.6909
16.805 | -35.307 | -30.6994
16.8111 | -35.3077 | -30.7077
16.8172 | -35.3084 | -30.7162
16.8233 | -35.3091 | -30.7247
16.8294 | -35.3098 | -30.7333
16.8356 | -35.3105 | -30.742
16.8417 | -35.3112 | -30.7503
16.8478 | -35.3119 | -30.7583
16.8539 | -35.3126 | -30.7668
16.86 | -35.3133 | -30.7752
16.8661 | -35.314 | -30.7835
16.8722 | -35.3148 | -30.7921
16.8783 | -35.3154 | -30.8
16.8844 | -35.3161 | -30.808
16.8905 | -35.3168 | -30.8164
16.8966 | -35.3175 | -30.8251
16.9027 | -35.3182 | -30.8332
16.9088 | -35.3189 | -30.8421
16.9149 | -35.3196 | -30.851
16.921 | -35.3203 | -30.859
16.9271 | -35.321 | -30.8677
16.9332 | -35.3217 | -30.8763
16.9393 | -35.3224 | -30.885
16.9454 | -35.323 | -30.8933
16.9515 | -35.3237 | -30.902
16.9576 | -35.3244 | -30.9103
16.9637 | -35.3251 | -30.9186
16.9698 | -35.3258 | -30.9267
16.9759 | -35.3265 | -30.9349
16.9821 | -35.3272 | -30.9434
16.9882 | -35.3279 | -30.9518
16.9943 | -35.3286 | -30.9603
17.0004 | -35.3293 | -30.9687
17.0065 | -35.33 | -30.9771
17.0126 | -35.3307 | -30.9853
17.0187 | -35.3313 | -30.9934
17.0248 | -35.332 | -31.0014
17.0309 | -35.3327 | -31.0096
17.037 | -35.3333 | -31.0181
17.0431 | -35.3341 | -31.026
17.0492 | -35.3348 | -31.0344
17.0553 | -35.3354 | -31.0431
17.0614 | -35.3361 | -31.0511
17.0675 | -35.3368 | -31.0595
17.0736 | -35.3375 | -31.0679
17.0797 | -35.3382 | -31.0759
17.0858 | -35.3389 | -31.0842
17.0919 | -35.3396 | -31.0927
17.098 | -35.3403 | -31.1008
17.1041 | -35.3409 | -31.1088
17.1102 | -35.3416 | -31.1173
17.1163 | -35.3423 | -31.1257
17.1225 | -35.343 | -31.1339
17.1286 | -35.3437 | -31.142
17.1347 | -35.3444 | -31.1503
17.1408 | -35.3451 | -31.1586
17.1469 | -35.3458 | -31.1669
17.153 | -35.3465 | -31.1751
17.1591 | -35.3472 | -31.183
17.1652 | -35.3479 | -31.191
17.1713 | -35.3486 | -31.1994
17.1774 | -35.3492 | -31.2078
17.1835 | -35.3499 | -31.2161
17.1896 | -35.3506 | -31.2245
17.1957 | -35.3513 | -31.233
17.2018 | -35.352 | -31.241
17.2079 | -35.3527 | -31.2492
17.214 | -35.3534 | -31.2575
17.2201 | -35.354 | -31.2659
17.2262 | -35.3547 | -31.2742
17.2323 | -35.3554 | -31.2822
17.2384 | -35.3562 | -31.2906
17.2445 | -35.3569 | -31.299
17.2506 | -35.3576 | -31.3072
17.2567 | -35.3583 | -31.3154
17.2628 | -35.359 | -31.3238
17.269 | -35.3597 | -31.3319
17.2751 | -35.3603 | -31.3402
17.2812 | -35.361 | -31.3483
17.2873 | -35.3617 | -31.3567
17.2934 | -35.3624 | -31.365
17.2995 | -35.363 | -31.373
17.3056 | -35.3637 | -31.3808
17.3117 | -35.3644 | -31.3892
17.3178 | -35.3651 | -31.3973
17.3239 | -35.3657 | -31.4054
17.33 | -35.3664 | -31.4133
17.3361 | -35.3671 | -31.4214
17.3422 | -35.3678 | -31.4297
17.3483 | -35.3684 | -31.438
17.3544 | -35.3691 | -31.4463
17.3605 | -35.3698 | -31.4545
17.3666 | -35.3705 | -31.4628
17.3727 | -35.3712 | -31.471
17.3788 | -35.3718 | -31.4788
17.3849 | -35.3725 | -31.487
17.391 | -35.3732 | -31.4952
17.3971 | -35.374 | -31.5036
17.4032 | -35.3746 | -31.5121
17.4094 | -35.3753 | -31.5198
17.4155 | -35.376 | -31.5279
17.4216 | -35.3767 | -31.536
17.4277 | -35.3773 | -31.5445
17.4338 | -35.3781 | -31.5526
17.4399 | -35.3788 | -31.5609
17.446 | -35.3794 | -31.5686
17.4521 | -35.3801 | -31.5768
17.4582 | -35.3808 | -31.5849
17.4643 | -35.3815 | -31.593
17.4704 | -35.3821 | -31.6011
17.4765 | -35.3828 | -31.6091
17.4826 | -35.3835 | -31.6174
17.4887 | -35.3842 | -31.6259
17.4948 | -35.3848 | -31.634
17.5009 | -35.3855 | -31.6419
17.507 | -35.3862 | -31.6501
17.5131 | -35.3869 | -31.6584
17.5192 | -35.3875 | -31.6664
17.5253 | -35.3882 | -31.6746
17.5314 | -35.3888 | -31.6828
17.5375 | -35.3895 | -31.6914
17.5436 | -35.3901 | -31.6997
17.5497 | -35.3908 | -31.7081
17.5559 | -35.3914 | -31.7161
17.562 | -35.3921 | -31.7242
17.5681 | -35.3928 | -31.7322
17.5742 | -35.3935 | -31.7403
17.5803 | -35.3942 | -31.7483
17.5864 | -35.3949 | -31.7561
17.5925 | -35.3957 | -31.7644
17.5986 | -35.3964 | -31.7724
17.6047 | -35.3971 | -31.781
17.6108 | -35.3977 | -31.7889
17.6169 | -35.3984 | -31.7968
17.623 | -35.3991 | -31.8044
17.6291 | -35.3998 | -31.8124
17.6352 | -35.4005 | -31.8203
17.6413 | -35.4012 | -31.8285
17.6474 | -35.4018 | -31.8365
17.6535 | -35.4025 | -31.8446
17.6596 | -35.4032 | -31.8527
17.6657 | -35.4038 | -31.8606
17.6718 | -35.4045 | -31.8687
17.6779 | -35.4051 | -31.8767
17.684 | -35.4058 | -31.8847
17.6901 | -35.4065 | -31.8928
17.6963 | -35.4072 | -31.9007
17.7024 | -35.4079 | -31.9087
17.7085 | -35.4085 | -31.9164
17.7146 | -35.4093 | -31.9244
17.7207 | -35.41 | -31.9327
17.7268 | -35.4107 | -31.9408
17.7329 | -35.4113 | -31.9491
17.739 | -35.412 | -31.9573
17.7451 | -35.4127 | -31.9653
17.7512 | -35.4134 | -31.9733
17.7573 | -35.414 | -31.9815
17.7634 | -35.4147 | -31.9891
17.7695 | -35.4153 | -31.9972
17.7756 | -35.416 | -32.0051
17.7817 | -35.4167 | -32.0131
17.7878 | -35.4174 | -32.0214
17.7939 | -35.4181 | -32.0294
17.8 | -35.4188 | -32.0371
17.8061 | -35.4194 | -32.0451
17.8122 | -35.4201 | -32.053
17.8183 | -35.4208 | -32.061
17.8244 | -35.4215 | -32.069
17.8305 | -35.4222 | -32.0771
17.8366 | -35.4229 | -32.0847
17.8428 | -35.4236 | -32.0929
17.8489 | -35.4242 | -32.101
17.855 | -35.4249 | -32.109
17.8611 | -35.4256 | -32.117
17.8672 | -35.4263 | -32.1249
17.8733 | -35.4269 | -32.133
17.8794 | -35.4277 | -32.1413
17.8855 | -35.4283 | -32.1492
17.8916 | -35.429 | -32.1569
17.8977 | -35.4298 | -32.1651
17.9038 | -35.4304 | -32.1724
17.9099 | -35.4311 | -32.1802
17.916 | -35.4318 | -32.1883
17.9221 | -35.4325 | -32.1962
17.9282 | -35.4332 | -32.2041
17.9343 | -35.4339 | -32.2124
17.9404 | -35.4346 | -32.22
17.9465 | -35.4352 | -32.2279
17.9526 | -35.4358 | -32.2361
17.9587 | -35.4365 | -32.2443
17.9648 | -35.4372 | -32.2521
17.9709 | -35.4379 | -32.2601
17.977 | -35.4386 | -32.2679
17.9832 | -35.4392 | -32.2758
17.9893 | -35.4399 | -32.2838
17.9954 | -35.4406 | -32.2919
18.0015 | -35.4413 | -32.2999
18.0076 | -35.442 | -32.3072
18.0137 | -35.4426 | -32.3156
18.0198 | -35.4433 | -32.3234
18.0259 | -35.4439 | -32.3315
18.032 | -35.4445 | -32.3393
18.0381 | -35.4452 | -32.3471
18.0442 | -35.4459 | -32.355
18.0503 | -35.4466 | -32.3632
18.0564 | -35.4473 | -32.371
18.0625 | -35.448 | -32.3792
18.0686 | -35.4487 | -32.3872
18.0747 | -35.4494 | -32.395
18.0808 | -35.4501 | -32.4027
18.0869 | -35.4508 | -32.4108
18.093 | -35.4515 | -32.419
18.0991 | -35.4522 | -32.427
18.1052 | -35.4528 | -32.4352
18.1113 | -35.4535 | -32.443
18.1174 | -35.4542 | -32.4505
18.1236 | -35.4549 | -32.4586
18.1297 | -35.4555 | -32.4663
18.1358 | -35.4563 | -32.4742
18.1419 | -35.457 | -32.4823
18.148 | -35.4577 | -32.4903
18.1541 | -35.4584 | -32.4978
18.1602 | -35.4591 | -32.506
18.1663 | -35.4599 | -32.5138
18.1724 | -35.4606 | -32.5218
18.1785 | -35.4612 | -32.53
18.1846 | -35.4619 | -32.5378
18.1907 | -35.4625 | -32.546
18.1968 | -35.4632 | -32.5543
18.2029 | -35.4639 | -32.562
18.209 | -35.4645 | -32.5701
18.2151 | -35.4653 | -32.5781
18.2212 | -35.466 | -32.5853
18.2273 | -35.4667 | -32.5929
18.2334 | -35.4674 | -32.6009
18.2395 | -35.468 | -32.6083
18.2456 | -35.4687 | -32.6158
18.2517 | -35.4693 | -32.6236
18.2578 | -35.47 | -32.6312
18.2639 | -35.4708 | -32.6393
18.2701 | -35.4714 | -32.6469
18.2762 | -35.4722 | -32.655
18.2823 | -35.4729 | -32.6627
18.2884 | -35.4736 | -32.6706
18.2945 | -35.4743 | -32.678
18.3006 | -35.475 | -32.6857
18.3067 | -35.4757 | -32.6932
18.3128 | -35.4764 | -32.7012
18.3189 | -35.477 | -32.709
18.325 | -35.4777 | -32.7166
18.3311 | -35.4784 | -32.7243
18.3372 | -35.479 | -32.7321
18.3433 | -35.4797 | -32.7398
18.3494 | -35.4803 | -32.7476
18.3555 | -35.481 | -32.7554
18.3616 | -35.4817 | -32.7626
18.3677 | -35.4824 | -32.7702
18.3738 | -35.4831 | -32.7779
18.3799 | -35.4837 | -32.7855
18.386 | -35.4844 | -32.7933
18.3921 | -35.4851 | -32.8009
18.3982 | -35.4857 | -32.8086
18.4043 | -35.4863 | -32.8162
18.4105 | -35.487 | -32.8239
18.4166 | -35.4877 | -32.8317
18.4227 | -35.4884 | -32.8397
18.4288 | -35.4891 | -32.8475
18.4349 | -35.4898 | -32.8548
18.441 | -35.4905 | -32.8627
18.4471 | -35.4912 | -32.8705
18.4532 | -35.4918 | -32.8783
18.4593 | -35.4925 | -32.8861
18.4654 | -35.4931 | -32.8933
18.4715 | -35.4938 | -32.9011
18.4776 | -35.4945 | -32.9088
18.4837 | -35.4952 | -32.9166
18.4898 | -35.4959 | -32.9247
18.4959 | -35.4966 | -32.9325
18.502 | -35.4973 | -32.9403
18.5081 | -35.498 | -32.9481
18.5142 | -35.4987 | -32.956
18.5203 | -35.4994 | -32.9635
18.5264 | -35.5 | -32.9709
18.5325 | -35.5007 | -32.9789
18.5386 | -35.5014 | -32.9867
18.5447 | -35.5021 | -32.9944
18.5508 | -35.5028 | -33.0019
18.557 | -35.5035 | -33.0098
18.5631 | -35.5041 | -33.0174
18.5692 | -35.5048 | -33.0249
18.5753 | -35.5055 | -33.0323
18.5814 | -35.5062 | -33.04
18.5875 | -35.5069 | -33.0481
18.5936 | -35.5076 | -33.0564
18.5997 | -35.5083 | -33.0645
18.6058 | -35.5089 | -33.0723
18.6119 | -35.5096 | -33.0802
18.618 | -35.5103 | -33.0882
18.6241 | -35.5109 | -33.0962
18.6302 | -35.5117 | -33.1038
18.6363 | -35.5124 | -33.1114
18.6424 | -35.513 | -33.119
18.6485 | -35.5137 | -33.1266
18.6546 | -35.5144 | -33.1342
18.6607 | -35.5151 | -33.1418
18.6668 | -35.5158 | -33.1496
18.6729 | -35.5164 | -33.1575
18.679 | -35.5171 | -33.1648
18.6851 | -35.5178 | -33.1721
18.6912 | -35.5185 | -33.1796
18.6974 | -35.5192 | -33.1873
18.7035 | -35.5199 | -33.1951
18.7096 | -35.5206 | -33.2028
18.7157 | -35.5213 | -33.2102
18.7218 | -35.522 | -33.218
18.7279 | -35.5227 | -33.2254
18.734 | -35.5234 | -33.233
18.7401 | -35.5241 | -33.2404
18.7462 | -35.5248 | -33.2473
18.7523 | -35.5255 | -33.2548
18.7584 | -35.5262 | -33.2628
18.7645 | -35.5269 | -33.2705
18.7706 | -35.5275 | -33.2782
18.7767 | -35.5282 | -33.2859
18.7828 | -35.5289 | -33.2936
18.7889 | -35.5296 | -33.3012
18.795 | -35.5302 | -33.3086
18.8011 | -35.531 | -33.3161
18.8072 | -35.5317 | -33.3236
18.8133 | -35.5323 | -33.3314
18.8194 | -35.5329 | -33.3383
18.8255 | -35.5337 | -33.3458
18.8316 | -35.5343 | -33.3538
18.8377 | -35.535 | -33.3613
18.8439 | -35.5357 | -33.3688
18.85 | -35.5364 | -33.3764
18.8561 | -35.5371 | -33.3839
18.8622 | -35.5378 | -33.3918
18.8683 | -35.5385 | -33.399
18.8744 | -35.5392 | -33.4065
18.8805 | -35.5399 | -33.4144
18.8866 | -35.5406 | -33.4222
18.8927 | -35.5413 | -33.4294
18.8988 | -35.542 | -33.437
18.9049 | -35.5427 | -33.4446
18.911 | -35.5434 | -33.4522
18.9171 | -35.5441 | -33.4598
18.9232 | -35.5448 | -33.467
18.9293 | -35.5455 | -33.4748
18.9354 | -35.5461 | -33.4822
18.9415 | -35.5468 | -33.4897
18.9476 | -35.5475 | -33.4974
18.9537 | -35.5482 | -33.5051
18.9598 | -35.5489 | -33.5123
18.9659 | -35.5496 | -33.52
18.972 | -35.5502 | -33.5274
18.9781 | -35.5509 | -33.5346
18.9843 | -35.5516 | -33.5424
18.9904 | -35.5522 | -33.55
18.9965 | -35.5529 | -33.5577
19.0026 | -35.5536 | -33.5651
19.0087 | -35.5542 | -33.5725
19.0148 | -35.555 | -33.5799
19.0209 | -35.5557 | -33.5877
19.027 | -35.5564 | -33.5955
19.0331 | -35.5571 | -33.603
19.0392 | -35.5578 | -33.6106
19.0453 | -35.5585 | -33.618
19.0514 | -35.5591 | -33.6257
19.0575 | -35.5598 | -33.6332
19.0636 | -35.5605 | -33.6403
19.0697 | -35.5612 | -33.6479
19.0758 | -35.562 | -33.6554
19.0819 | -35.5626 | -33.663
19.088 | -35.5633 | -33.6704
19.0941 | -35.564 | -33.6774
19.1002 | -35.5646 | -33.6838
19.1063 | -35.5653 | -33.6914
19.1124 | -35.566 | -33.6988
19.1185 | -35.5667 | -33.7064
19.1246 | -35.5674 | -33.7136
19.1308 | -35.5682 | -33.721
19.1369 | -35.5689 | -33.7282
19.143 | -35.5695 | -33.7359
19.1491 | -35.5703 | -33.7433
19.1552 | -35.571 | -33.7513
19.1613 | -35.5716 | -33.7588
19.1674 | -35.5723 | -33.7665
19.1735 | -35.573 | -33.7739
19.1796 | -35.5737 | -33.7812
19.1857 | -35.5744 | -33.7884
19.1918 | -35.575 | -33.7959
19.1979 | -35.5756 | -33.8033
19.204 | -35.5763 | -33.8105
19.2101 | -35.577 | -33.818
19.2162 | -35.5777 | -33.8258
19.2223 | -35.5784 | -33.8331
19.2284 | -35.5791 | -33.8405
19.2345 | -35.5798 | -33.8477
19.2406 | -35.5805 | -33.855
19.2467 | -35.5812 | -33.8624
19.2528 | -35.5818 | -33.8699
19.2589 | -35.5825 | -33.8773
19.265 | -35.5832 | -33.8848
19.2712 | -35.5838 | -33.8921
19.2773 | -35.5845 | -33.8994
19.2834 | -35.5852 | -33.9066
19.2895 | -35.5859 | -33.914
19.2956 | -35.5866 | -33.9218
19.3017 | -35.5873 | -33.9291
19.3078 | -35.5879 | -33.9365
19.3139 | -35.5887 | -33.9438
19.32 | -35.5894 | -33.9508
19.3261 | -35.59 | -33.9584
19.3322 | -35.5908 | -33.9657
19.3383 | -35.5915 | -33.9733
19.3444 | -35.5922 | -33.9805
19.3505 | -35.5928 | -33.9881
19.3566 | -35.5935 | -33.9956
19.3627 | -35.5942 | -34.0032
19.3688 | -35.5949 | -34.0103
19.3749 | -35.5956 | -34.0178
19.381 | -35.5964 | -34.0246
19.3871 | -35.597 | -34.0319
19.3932 | -35.5977 | -34.0395
19.3993 | -35.5984 | -34.0467
19.4054 | -35.5991 | -34.0545
19.4115 | -35.5997 | -34.0618
19.4177 | -35.6003 | -34.0689
19.4238 | -35.601 | -34.076
19.4299 | -35.6018 | -34.0835
19.436 | -35.6024 | -34.0911
19.4421 | -35.6031 | -34.0981
19.4482 | -35.6038 | -34.1056
19.4543 | -35.6045 | -34.1125
19.4604 | -35.6052 | -34.1195
19.4665 | -35.606 | -34.1269
19.4726 | -35.6067 | -34.1342
19.4787 | -35.6073 | -34.1416
19.4848 | -35.608 | -34.149
19.4909 | -35.6086 | -34.1561
19.497 | -35.6092 | -34.1633
19.5031 | -35.6099 | -34.1705
19.5092 | -35.6106 | -34.1779
19.5153 | -35.6113 | -34.1854
19.5214 | -35.612 | -34.1923
19.5275 | -35.6126 | -34.1991
19.5336 | -35.6132 | -34.2063
19.5397 | -35.6139 | -34.2136
19.5458 | -35.6146 | -34.2212
19.5519 | -35.6153 | -34.2285
19.5581 | -35.616 | -34.2353
19.5642 | -35.6167 | -34.2425
19.5703 | -35.6173 | -34.25
19.5764 | -35.618 | -34.2575
19.5825 | -35.6186 | -34.2649
19.5886 | -35.6193 | -34.2718
19.5947 | -35.62 | -34.2791
19.6008 | -35.6206 | -34.2865
19.6069 | -35.6213 | -34.2939
19.613 | -35.622 | -34.3014
19.6191 | -35.6227 | -34.3088
19.6252 | -35.6234 | -34.3162
19.6313 | -35.6241 | -34.3238
19.6374 | -35.6248 | -34.3314
19.6435 | -35.6255 | -34.3382
19.6496 | -35.6262 | -34.3455
19.6557 | -35.6269 | -34.3531
19.6618 | -35.6275 | -34.3599
19.6679 | -35.6282 | -34.3671
19.674 | -35.6289 | -34.3743
19.6801 | -35.6295 | -34.3817
19.6862 | -35.6302 | -34.3893
19.6923 | -35.631 | -34.3965
19.6984 | -35.6316 | -34.4034
19.7046 | -35.6324 | -34.4105
19.7107 | -35.6331 | -34.4177
19.7168 | -35.6337 | -34.4253
19.7229 | -35.6343 | -34.4327
19.729 | -35.635 | -34.4403
19.7351 | -35.6356 | -34.4475
19.7412 | -35.6363 | -34.4549
19.7473 | -35.637 | -34.4622
19.7534 | -35.6377 | -34.4699
19.7595 | -35.6383 | -34.4774
19.7656 | -35.639 | -34.4843
19.7717 | -35.6397 | -34.4914
19.7778 | -35.6404 | -34.4986
19.7839 | -35.6411 | -34.5054
19.79 | -35.6418 | -34.5124
19.7961 | -35.6423 | -34.5201
19.8022 | -35.643 | -34.5272
19.8083 | -35.6437 | -34.5342
19.8144 | -35.6443 | -34.541
19.8205 | -35.6449 | -34.5485
19.8266 | -35.6456 | -34.5555
19.8327 | -35.6462 | -34.5627
19.8388 | -35.6469 | -34.57
19.845 | -35.6476 | -34.5771
19.8511 | -35.6482 | -34.5846
19.8572 | -35.6489 | -34.5921
19.8633 | -35.6496 | -34.599
19.8694 | -35.6503 | -34.6062
19.8755 | -35.651 | -34.6129
19.8816 | -35.6517 | -34.6203
19.8877 | -35.6523 | -34.6278
19.8938 | -35.653 | -34.6351
19.8999 | -35.6537 | -34.6423
19.906 | -35.6544 | -34.65
19.9121 | -35.655 | -34.6567
19.9182 | -35.6558 | -34.6638
19.9243 | -35.6564 | -34.671
19.9304 | -35.6571 | -34.6784
19.9365 | -35.6578 | -34.6858
19.9426 | -35.6584 | -34.6928
19.9487 | -35.6591 | -34.6998
19.9548 | -35.6597 | -34.7069
19.9609 | -35.6604 | -34.7139
19.967 | -35.6611 | -34.721
19.9731 | -35.6617 | -34.7283
19.9792 | -35.6624 | -34.7352
19.9853 | -35.6631 | -34.7424
19.9915 | -35.6637 | -34.7498
19.9976 | -35.6644 | -34.7568
20.0037 | -35.6651 | -34.7638
20.0098 | -35.6657 | -34.7709
20.0159 | -35.6664 | -34.7777
20.022 | -35.667 | -34.7851
20.0281 | -35.6677 | -34.7926
20.0342 | -35.6684 | -34.7996
20.0403 | -35.6691 | -34.8064
20.0464 | -35.6697 | -34.8139
20.0525 | -35.6705 | -34.8211
20.0586 | -35.6711 | -34.8284
20.0647 | -35.6718 | -34.8357
20.0708 | -35.6725 | -34.8432
20.0769 | -35.6731 | -34.8504
20.083 | -35.6738 | -34.8574
20.0891 | -35.6744 | -34.8646
20.0952 | -35.6751 | -34.8715
20.1013 | -35.6757 | -34.8784
20.1074 | -35.6763 | -34.886
20.1135 | -35.677 | -34.8934
20.1196 | -35.6776 | -34.9001
20.1257 | -35.6783 | -34.9073
20.1319 | -35.6789 | -34.9142
20.138 | -35.6796 | -34.9211
20.1441 | -35.6802 | -34.928
20.1502 | -35.6809 | -34.9355
20.1563 | -35.6816 | -34.9421
20.1624 | -35.6823 | -34.9493
20.1685 | -35.683 | -34.9565
20.1746 | -35.6835 | -34.9636
20.1807 | -35.6842 | -34.9707
20.1868 | -35.6848 | -34.9779
20.1929 | -35.6855 | -34.9849
20.199 | -35.6861 | -34.9923
20.2051 | -35.6868 | -34.9993
20.2112 | -35.6874 | -35.0064
20.2173 | -35.6881 | -35.0135
20.2234 | -35.6887 | -35.0207
20.2295 | -35.6894 | -35.0276
20.2356 | -35.6901 | -35.035
20.2417 | -35.6907 | -35.0421
20.2478 | -35.6913 | -35.049
20.2539 | -35.6919 | -35.0563
20.26 | -35.6925 | -35.0634
20.2661 | -35.6932 | -35.0706
20.2723 | -35.6938 | -35.0779
20.2784 | -35.6945 | -35.085
20.2845 | -35.6951 | -35.0921
20.2906 | -35.6958 | -35.0995
20.2967 | -35.6964 | -35.1066
20.3028 | -35.6971 | -35.1136
20.3089 | -35.6977 | -35.1206
20.315 | -35.6984 | -35.1278
20.3211 | -35.699 | -35.1349
20.3272 | -35.6997 | -35.1419
20.3333 | -35.7004 | -35.1488
20.3394 | -35.701 | -35.1558
20.3455 | -35.7017 | -35.1629
20.3516 | -35.7023 | -35.1699
20.3577 | -35.703 | -35.1767
20.3638 | -35.7037 | -35.1837
20.3699 | -35.7043 | -35.1906
20.376 | -35.705 | -35.1977
20.3821 | -35.7056 | -35.2049
20.3882 | -35.7063 | -35.212
20.3943 | -35.707 | -35.2192
20.4004 | -35.7076 | -35.2258
20.4065 | -35.7082 | -35.2328
20.4126 | -35.7089 | -35.2401
20.4188 | -35.7095 | -35.2473
20.4249 | -35.7102 | -35.2547
20.431 | -35.7108 | -35.2617
20.4371 | -35.7115 | -35.2682
20.4432 | -35.7121 | -35.2755
20.4493 | -35.7127 | -35.2823
20.4554 | -35.7134 | -35.2894
20.4615 | -35.714 | -35.2966
20.4676 | -35.7146 | -35.3036
20.4737 | -35.7152 | -35.3103
20.4798 | -35.7159 | -35.3171
20.4859 | -35.7166 | -35.3244
20.492 | -35.7172 | -35.3314
20.4981 | -35.7179 | -35.3381
20.5042 | -35.7185 | -35.3451
20.5103 | -35.7192 | -35.3518
20.5164 | -35.7198 | -35.3585
20.5225 | -35.7204 | -35.3658
20.5286 | -35.7209 | -35.3732
20.5347 | -35.7216 | -35.3803
20.5408 | -35.7222 | -35.387
20.5469 | -35.7228 | -35.394
20.553 | -35.7235 | -35.401
20.5592 | -35.7241 | -35.4081
20.5653 | -35.7248 | -35.415
20.5714 | -35.7254 | -35.422
20.5775 | -35.726 | -35.4287
20.5836 | -35.7266 | -35.4356
20.5897 | -35.7273 | -35.4426
20.5958 | -35.7279 | -35.4497
20.6019 | -35.7285 | -35.4568
20.608 | -35.7292 | -35.4638
20.6141 | -35.7298 | -35.471
20.6202 | -35.7305 | -35.4778
20.6263 | -35.7311 | -35.4848
20.6324 | -35.7318 | -35.492
20.6385 | -35.7324 | -35.4988
20.6446 | -35.733 | -35.5061
20.6507 | -35.7336 | -35.5133
20.6568 | -35.7343 | -35.5205
20.6629 | -35.735 | -35.5274
20.669 | -35.7356 | -35.5346
20.6751 | -35.7362 | -35.5417
20.6812 | -35.7368 | -35.5484
20.6873 | -35.7375 | -35.5555
20.6934 | -35.7382 | -35.5627
20.6995 | -35.7388 | -35.5696
20.7057 | -35.7394 | -35.5768
20.7118 | -35.74 | -35.5836
20.7179 | -35.7406 | -35.59
20.724 | -35.7412 | -35.5966
20.7301 | -35.7418 | -35.6038
20.7362 | -35.7424 | -35.6109
20.7423 | -35.7431 | -35.6178
20.7484 | -35.7437 | -35.625
20.7545 | -35.7444 | -35.6318
20.7606 | -35.7451 | -35.639
20.7667 | -35.7457 | -35.6462
20.7728 | -35.7463 | -35.6531
20.7789 | -35.747 | -35.66
20.785 | -35.7476 | -35.667
20.7911 | -35.7482 | -35.674
20.7972 | -35.7488 | -35.681
20.8033 | -35.7494 | -35.6876
20.8094 | -35.7501 | -35.6946
20.8155 | -35.7508 | -35.7014
20.8216 | -35.7514 | -35.7083
20.8277 | -35.752 | -35.715
20.8338 | -35.7526 | -35.7221
20.8399 | -35.7532 | -35.7289
20.8461 | -35.7539 | -35.7352
20.8522 | -35.7545 | -35.742
20.8583 | -35.7551 | -35.7488
20.8644 | -35.7558 | -35.7556
20.8705 | -35.7564 | -35.7626
20.8766 | -35.7571 | -35.7696
20.8827 | -35.7578 | -35.7767
20.8888 | -35.7584 | -35.7835
20.8949 | -35.759 | -35.7906
20.901 | -35.7596 | -35.7973
20.9071 | -35.7602 | -35.8041
20.9132 | -35.7608 | -35.8111
20.9193 | -35.7615 | -35.818
20.9254 | -35.7622 | -35.8248
20.9315 | -35.7628 | -35.831
20.9376 | -35.7633 | -35.8381
20.9437 | -35.764 | -35.845
20.9498 | -35.7646 | -35.8521
20.9559 | -35.7653 | -35.8586
20.962 | -35.766 | -35.8658
20.9681 | -35.7665 | -35.8726
20.9742 | -35.7672 | -35.8793
20.9803 | -35.7678 | -35.8864
20.9864 | -35.7684 | -35.8933
20.9926 | -35.7691 | -35.9004
20.9987 | -35.7697 | -35.9067
21.0048 | -35.7703 | -35.9137
21.0109 | -35.7709 | -35.9204
21.017 | -35.7715 | -35.9273
21.0231 | -35.7722 | -35.9339
21.0292 | -35.7728 | -35.9409
21.0353 | -35.7735 | -35.9474
21.0414 | -35.7741 | -35.9545
21.0475 | -35.7747 | -35.9615
21.0536 | -35.7753 | -35.9683
21.0597 | -35.776 | -35.9749
21.0658 | -35.7766 | -35.9815
21.0719 | -35.7772 | -35.9889
21.078 | -35.7778 | -35.9962
21.0841 | -35.7784 | -36.0027
21.0902 | -35.779 | -36.0096
21.0963 | -35.7797 | -36.0164
21.1024 | -35.7803 | -36.0235
21.1085 | -35.7809 | -36.0304
21.1146 | -35.7815 | -36.0371
21.1207 | -35.782 | -36.0441
21.1268 | -35.7827 | -36.0503
21.133 | -35.7833 | -36.057
21.1391 | -35.784 | -36.0636
21.1452 | -35.7846 | -36.0706
21.1513 | -35.7852 | -36.0773
21.1574 | -35.7858 | -36.0843
21.1635 | -35.7864 | -36.0912
21.1696 | -35.787 | -36.0978
21.1757 | -35.7876 | -36.1039
21.1818 | -35.7883 | -36.111
21.1879 | -35.7889 | -36.1179
21.194 | -35.7896 | -36.1248
21.2001 | -35.7902 | -36.1313
21.2062 | -35.7908 | -36.1378
21.2123 | -35.7914 | -36.1445
21.2184 | -35.792 | -36.1514
21.2245 | -35.7927 | -36.1582
21.2306 | -35.7933 | -36.1647
21.2367 | -35.7939 | -36.1719
21.2428 | -35.7945 | -36.1787
21.2489 | -35.7951 | -36.1852
21.255 | -35.7958 | -36.1919
21.2611 | -35.7964 | -36.1988
21.2672 | -35.7969 | -36.2056
21.2733 | -35.7975 | -36.2122
21.2795 | -35.7982 | -36.2188
21.2856 | -35.7988 | -36.2256
21.2917 | -35.7995 | -36.2323
21.2978 | -35.8001 | -36.239
21.3039 | -35.8007 | -36.2457
21.31 | -35.8013 | -36.2524
21.3161 | -35.8019 | -36.2589
21.3222 | -35.8025 | -36.2654
21.3283 | -35.8032 | -36.2719
21.3344 | -35.8037 | -36.2784
21.3405 | -35.8043 | -36.2854
21.3466 | -35.805 | -36.2921
21.3527 | -35.8056 | -36.2987
21.3588 | -35.8063 | -36.3056
21.3649 | -35.8068 | -36.312
21.371 | -35.8075 | -36.3181
21.3771 | -35.8082 | -36.3247
21.3832 | -35.8088 | -36.3311
21.3893 | -35.8094 | -36.3377
21.3954 | -35.81 | -36.3445
21.4015 | -35.8106 | -36.3512
21.4076 | -35.8113 | -36.3581
21.4137 | -35.8119 | -36.3652
21.4199 | -35.8124 | -36.372
21.426 | -35.813 | -36.3784
21.4321 | -35.8137 | -36.3854
21.4382 | -35.8143 | -36.3922
21.4443 | -35.8149 | -36.3991
21.4504 | -35.8155 | -36.4057
21.4565 | -35.8161 | -36.4123
21.4626 | -35.8168 | -36.4193
21.4687 | -35.8174 | -36.4261
21.4748 | -35.818 | -36.4324
21.4809 | -35.8185 | -36.4393
21.487 | -35.819 | -36.4457
21.4931 | -35.8196 | -36.4519
21.4992 | -35.8203 | -36.4581
21.5053 | -35.8208 | -36.4653
21.5114 | -35.8214 | -36.4718
21.5175 | -35.822 | -36.479
21.5236 | -35.8225 | -36.486
21.5297 | -35.8231 | -36.4928
21.5358 | -35.8238 | -36.4993
21.5419 | -35.8244 | -36.5063
21.548 | -35.825 | -36.513
21.5541 | -35.8256 | -36.5196
21.5602 | -35.8262 | -36.5261
21.5664 | -35.8269 | -36.5327
21.5725 | -35.8275 | -36.5392
21.5786 | -35.828 | -36.5456
21.5847 | -35.8286 | -36.5522
21.5908 | -35.8293 | -36.5592
21.5969 | -35.8299 | -36.5656
21.603 | -35.8305 | -36.5722
21.6091 | -35.8311 | -36.5791
21.6152 | -35.8317 | -36.5863
21.6213 | -35.8323 | -36.593
21.6274 | -35.8329 | -36.5996
21.6335 | -35.8335 | -36.606
21.6396 | -35.8341 | -36.6128
21.6457 | -35.8347 | -36.6197
21.6518 | -35.8353 | -36.626
21.6579 | -35.836 | -36.6323
21.664 | -35.8366 | -36.6393
21.6701 | -35.8372 | -36.6461
21.6762 | -35.8378 | -36.6528
21.6823 | -35.8383 | -36.6593
21.6884 | -35.8389 | -36.6658
21.6945 | -35.8395 | -36.6725
21.7006 | -35.8401 | -36.6793
21.7068 | -35.8407 | -36.6859
21.7129 | -35.8413 | -36.6925
21.719 | -35.8419 | -36.6994
21.7251 | -35.8425 | -36.7058
21.7312 | -35.8432 | -36.7121
21.7373 | -35.8438 | -36.719
21.7434 | -35.8444 | -36.7259
21.7495 | -35.845 | -36.7327
21.7556 | -35.8456 | -36.7398
21.7617 | -35.8462 | -36.7465
21.7678 | -35.8469 | -36.7535
21.7739 | -35.8474 | -36.76
21.78 | -35.848 | -36.7668
21.7861 | -35.8486 | -36.7734
21.7922 | -35.8492 | -36.7802
21.7983 | -35.8498 | -36.787
21.8044 | -35.8505 | -36.7938
21.8105 | -35.8511 | -36.8003
21.8166 | -35.8517 | -36.8068
21.8227 | -35.8523 | -36.8133
21.8288 | -35.8529 | -36.8199
21.8349 | -35.8535 | -36.8264
21.841 | -35.8541 | -36.8331
21.8471 | -35.8547 | -36.8392
21.8533 | -35.8553 | -36.8456
21.8594 | -35.8559 | -36.8524
21.8655 | -35.8564 | -36.8597
21.8716 | -35.8571 | -36.8664
21.8777 | -35.8576 | -36.873
21.8838 | -35.8582 | -36.8795
21.8899 | -35.8588 | -36.8859
21.896 | -35.8594 | -36.8927
21.9021 | -35.8601 | -36.8993
21.9082 | -35.8606 | -36.9058
21.9143 | -35.8612 | -36.9124
21.9204 | -35.8618 | -36.9191
21.9265 | -35.8623 | -36.9254
21.9326 | -35.863 | -36.9323
21.9387 | -35.8636 | -36.9387
21.9448 | -35.8642 | -36.9454
21.9509 | -35.8649 | -36.9516
21.957 | -35.8655 | -36.9579
21.9631 | -35.8661 | -36.9647
21.9692 | -35.8667 | -36.9713
21.9753 | -35.8673 | -36.9777
21.9814 | -35.8679 | -36.9845
21.9875 | -35.8685 | -36.9909
21.9937 | -35.8691 | -36.9977
21.9998 | -35.8697 | -37.0045
22.0059 | -35.8703 | -37.0112
22.012 | -35.8709 | -37.018
22.0181 | -35.8716 | -37.0238
22.0242 | -35.8722 | -37.0303
22.0303 | -35.8727 | -37.037
22.0364 | -35.8733 | -37.0434
22.0425 | -35.8739 | -37.0495
22.0486 | -35.8745 | -37.0558
22.0547 | -35.8751 | -37.0621
22.0608 | -35.8756 | -37.0686
22.0669 | -35.8762 | -37.0751
22.073 | -35.8768 | -37.0814
22.0791 | -35.8775 | -37.0881
22.0852 | -35.8781 | -37.0947
22.0913 | -35.8787 | -37.1011
22.0974 | -35.8793 | -37.1074
22.1035 | -35.8799 | -37.114
22.1096 | -35.8805 | -37.1207
22.1157 | -35.8811 | -37.1273
22.1218 | -35.8816 | -37.134
22.1279 | -35.8822 | -37.1401
22.134 | -35.8828 | -37.1467
22.1402 | -35.8834 | -37.1533
22.1463 | -35.8841 | -37.1597
22.1524 | -35.8847 | -37.1663
22.1585 | -35.8852 | -37.1731
22.1646 | -35.8859 | -37.1792
22.1707 | -35.8865 | -37.1861
22.1768 | -35.8871 | -37.1923
22.1829 | -35.8877 | -37.1988
22.189 | -35.8883 | -37.2055
22.1951 | -35.8889 | -37.2119
22.2012 | -35.8895 | -37.2185
22.2073 | -35.89 | -37.2253
22.2134 | -35.8906 | -37.2319
22.2195 | -35.8912 | -37.2384
22.2256 | -35.8918 | -37.2446
22.2317 | -35.8924 | -37.2507
22.2378 | -35.893 | -37.257
22.2439 | -35.8937 | -37.2634
22.25 | -35.8943 | -37.27
22.2561 | -35.8949 | -37.2767
22.2622 | -35.8955 | -37.2832
22.2683 | -35.8961 | -37.2894
22.2744 | -35.8967 | -37.2957
22.2806 | -35.8974 | -37.3024
22.2867 | -35.898 | -37.309
22.2928 | -35.8986 | -37.3154
22.2989 | -35.8992 | -37.3217
22.305 | -35.8998 | -37.3278
22.3111 | -35.9004 | -37.3347
22.3172 | -35.9009 | -37.3413
22.3233 | -35.9015 | -37.3479
22.3294 | -35.9021 | -37.3539
22.3355 | -35.9028 | -37.3602
22.3416 | -35.9034 | -37.3667
22.3477 | -35.904 | -37.3732
22.3538 | -35.9046 | -37.3793
22.3599 | -35.9052 | -37.3857
22.366 | -35.9058 | -37.3924
22.3721 | -35.9064 | -37.3988
22.3782 | -35.907 | -37.4056
22.3843 | -35.9076 | -37.4123
22.3904 | -35.9082 | -37.4191
22.3965 | -35.9088 | -37.4255
22.4026 | -35.9094 | -37.432
22.4087 | -35.91 | -37.4383
22.4148 | -35.9106 | -37.4448
22.4209 | -35.9112 | -37.4513
22.4271 | -35.9118 | -37.4576
22.4332 | -35.9124 | -37.4639
22.4393 | -35.913 | -37.4702
22.4454 | -35.9136 | -37.476
22.4515 | -35.9142 | -37.4825
22.4576 | -35.9148 | -37.4892
22.4637 | -35.9154 | -37.4955
22.4698 | -35.916 | -37.5019
22.4759 | -35.9166 | -37.5083
22.482 | -35.9172 | -37.515
22.4881 | -35.9178 | -37.5216
22.4942 | -35.9184 | -37.528
22.5003 | -35.919 | -37.5343
22.5064 | -35.9196 | -37.5406
22.5125 | -35.9201 | -37.547
22.5186 | -35.9208 | -37.5532
22.5247 | -35.9213 | -37.5597
22.5308 | -35.9219 | -37.5662
22.5369 | -35.9225 | -37.5727
22.543 | -35.9231 | -37.5789
22.5491 | -35.9237 | -37.5849
22.5552 | -35.9243 | -37.5914
22.5613 | -35.9249 | -37.5979
22.5675 | -35.9255 | -37.6046
22.5736 | -35.926 | -37.6112
22.5797 | -35.9266 | -37.6176
22.5858 | -35.9272 | -37.6239
22.5919 | -35.9277 | -37.6302
22.598 | -35.9284 | -37.6367
22.6041 | -35.929 | -37.6434
22.6102 | -35.9296 | -37.6498
22.6163 | -35.9301 | -37.656
22.6224 | -35.9307 | -37.6623
22.6285 | -35.9313 | -37.6686
22.6346 | -35.9319 | -37.6749
22.6407 | -35.9324 | -37.6811
22.6468 | -35.933 | -37.6875
22.6529 | -35.9336 | -37.6934
22.659 | -35.9342 | -37.6997
22.6651 | -35.9348 | -37.7063
22.6712 | -35.9353 | -37.7125
22.6773 | -35.9359 | -37.7187
22.6834 | -35.9365 | -37.7253
22.6895 | -35.9371 | -37.7317
22.6956 | -35.9378 | -37.7383
22.7017 | -35.9383 | -37.7446
22.7079 | -35.939 | -37.7509
22.714 | -35.9396 | -37.7572
22.7201 | -35.9402 | -37.7635
22.7262 | -35.9408 | -37.7696
22.7323 | -35.9414 | -37.776
22.7384 | -35.9419 | -37.7822
22.7445 | -35.9425 | -37.7888
22.7506 | -35.9431 | -37.795
22.7567 | -35.9437 | -37.8014
22.7628 | -35.9443 | -37.8073
22.7689 | -35.9449 | -37.8134
22.775 | -35.9455 | -37.8199
22.7811 | -35.9461 | -37.8264
22.7872 | -35.9467 | -37.8327
22.7933 | -35.9473 | -37.8389
22.7994 | -35.9479 | -37.8452
22.8055 | -35.9484 | -37.8516
22.8116 | -35.949 | -37.8579
22.8177 | -35.9496 | -37.8645
22.8238 | -35.9501 | -37.8712
22.8299 | -35.9507 | -37.8773
22.836 | -35.9514 | -37.884
22.8421 | -35.952 | -37.8903
22.8482 | -35.9525 | -37.8967
22.8544 | -35.9531 | -37.903
22.8605 | -35.9538 | -37.9092
22.8666 | -35.9543 | -37.9154
22.8727 | -35.9549 | -37.9216
22.8788 | -35.9556 | -37.9281
22.8849 | -35.9561 | -37.9344
22.891 | -35.9567 | -37.9409
22.8971 | -35.9573 | -37.9473
22.9032 | -35.9579 | -37.9532
22.9093 | -35.9584 | -37.9594
22.9154 | -35.959 | -37.9658
22.9215 | -35.9596 | -37.9721
22.9276 | -35.9602 | -37.9782
22.9337 | -35.9608 | -37.9843
22.9398 | -35.9613 | -37.9908
22.9459 | -35.962 | -37.9969
22.952 | -35.9625 | -38.0036
22.9581 | -35.9631 | -38.0099
22.9642 | -35.9637 | -38.0165
22.9703 | -35.9643 | -38.0223
22.9764 | -35.9649 | -38.0287
22.9825 | -35.9654 | -38.035
22.9886 | -35.966 | -38.0412
22.9948 | -35.9666 | -38.0476
23.0009 | -35.9672 | -38.0538
23.007 | -35.9678 | -38.0604
23.0131 | -35.9684 | -38.0665
23.0192 | -35.969 | -38.0725
23.0253 | -35.9696 | -38.0789
23.0314 | -35.9702 | -38.0853
23.0375 | -35.9708 | -38.0916
23.0436 | -35.9714 | -38.0976
23.0497 | -35.972 | -38.1043
23.0558 | -35.9726 | -38.1106
23.0619 | -35.9732 | -38.1169
23.068 | -35.9738 | -38.1235
23.0741 | -35.9743 | -38.1295
23.0802 | -35.9749 | -38.1356
23.0863 | -35.9755 | -38.142
23.0924 | -35.9761 | -38.1481
23.0985 | -35.9766 | -38.1545
23.1046 | -35.9772 | -38.1608
23.1107 | -35.9778 | -38.1673
23.1168 | -35.9784 | -38.1739
23.1229 | -35.979 | -38.1802
23.129 | -35.9796 | -38.1864
23.1351 | -35.9802 | -38.1925
23.1413 | -35.9808 | -38.1986
23.1474 | -35.9813 | -38.2046
23.1535 | -35.9819 | -38.2106
23.1596 | -35.9825 | -38.2167
23.1657 | -35.9831 | -38.223
23.1718 | -35.9837 | -38.2295
23.1779 | -35.9843 | -38.236
23.184 | -35.9849 | -38.2419
23.1901 | -35.9856 | -38.2482
23.1962 | -35.9861 | -38.2546
23.2023 | -35.9867 | -38.2607
23.2084 | -35.9873 | -38.2671
23.2145 | -35.9879 | -38.2735
23.2206 | -35.9885 | -38.2797
23.2267 | -35.989 | -38.2862
23.2328 | -35.9896 | -38.2926
23.2389 | -35.9902 | -38.2989
23.245 | -35.9908 | -38.3051
23.2511 | -35.9913 | -38.3111
23.2572 | -35.9919 | -38.3174
23.2633 | -35.9925 | -38.3237
23.2694 | -35.9931 | -38.3303
23.2755 | -35.9937 | -38.3367
23.2817 | -35.9943 | -38.3428
23.2878 | -35.9949 | -38.3492
23.2939 | -35.9955 | -38.3553
23.3 | -35.9961 | -38.3616
23.3061 | -35.9967 | -38.3677
23.3122 | -35.9973 | -38.3739
23.3183 | -35.9979 | -38.3802
23.3244 | -35.9984 | -38.3863
23.3305 | -35.999 | -38.3927
23.3366 | -35.9995 | -38.3993
23.3427 | -36.0001 | -38.4053
23.3488 | -36.0007 | -38.4115
23.3549 | -36.0013 | -38.4179
23.361 | -36.0019 | -38.424
23.3671 | -36.0025 | -38.4303
23.3732 | -36.0031 | -38.4364
23.3793 | -36.0037 | -38.4424
23.3854 | -36.0042 | -38.4489
23.3915 | -36.0048 | -38.4554
23.3976 | -36.0054 | -38.4617
23.4037 | -36.006 | -38.4677
23.4098 | -36.0066 | -38.4738
23.4159 | -36.0072 | -38.4801
23.422 | -36.0079 | -38.4864
23.4282 | -36.0085 | -38.4921
23.4343 | -36.0091 | -38.4982
23.4404 | -36.0096 | -38.5046
23.4465 | -36.0102 | -38.5105
23.4526 | -36.0108 | -38.5169
23.4587 | -36.0114 | -38.5232
23.4648 | -36.0119 | -38.529
23.4709 | -36.0125 | -38.5353
23.477 | -36.0131 | -38.5414
23.4831 | -36.0137 | -38.5475
23.4892 | -36.0143 | -38.554
23.4953 | -36.0149 | -38.5604
23.5014 | -36.0154 | -38.5665
23.5075 | -36.016 | -38.5728
23.5136 | -36.0166 | -38.5788
23.5197 | -36.0172 | -38.5852
23.5258 | -36.0177 | -38.5917
23.5319 | -36.0183 | -38.5979
23.538 | -36.0188 | -38.6041
23.5441 | -36.0194 | -38.6101
23.5502 | -36.0199 | -38.6167
23.5563 | -36.0205 | -38.6229
23.5624 | -36.0211 | -38.6295
23.5686 | -36.0217 | -38.6359
23.5747 | -36.0223 | -38.6419
23.5808 | -36.0228 | -38.6483
23.5869 | -36.0234 | -38.6545
23.593 | -36.024 | -38.6608
23.5991 | -36.0246 | -38.6672
23.6052 | -36.0252 | -38.6732
23.6113 | -36.0258 | -38.6796
23.6174 | -36.0263 | -38.6857
23.6235 | -36.0269 | -38.6916
23.6296 | -36.0275 | -38.6982
23.6357 | -36.0281 | -38.7047
23.6418 | -36.0286 | -38.7106
23.6479 | -36.0292 | -38.7169
23.654 | -36.0298 | -38.7228
23.6601 | -36.0304 | -38.7288
23.6662 | -36.031 | -38.7351
23.6723 | -36.0316 | -38.7409
23.6784 | -36.0322 | -38.7469
23.6845 | -36.0328 | -38.7531
23.6906 | -36.0333 | -38.7594
23.6967 | -36.0339 | -38.7657
23.7028 | -36.0345 | -38.7721
23.7089 | -36.035 | -38.7782
23.7151 | -36.0356 | -38.7843
23.7212 | -36.0362 | -38.7904
23.7273 | -36.0368 | -38.7966
23.7334 | -36.0375 | -38.8029
23.7395 | -36.038 | -38.8089
23.7456 | -36.0386 | -38.8148
23.7517 | -36.0392 | -38.8212
23.7578 | -36.0397 | -38.827
23.7639 | -36.0403 | -38.833
23.77 | -36.0409 | -38.8392
23.7761 | -36.0415 | -38.8453
23.7822 | -36.042 | -38.8516
23.7883 | -36.0426 | -38.8581
23.7944 | -36.0432 | -38.8644
23.8005 | -36.0438 | -38.8706
23.8066 | -36.0444 | -38.8768
23.8127 | -36.045 | -38.8828
23.8188 | -36.0455 | -38.8887
23.8249 | -36.0462 | -38.895
23.831 | -36.0467 | -38.9014
23.8371 | -36.0473 | -38.9075
23.8432 | -36.0478 | -38.9138
23.8493 | -36.0484 | -38.9195
23.8555 | -36.049 | -38.9257
23.8616 | -36.0496 | -38.9321
23.8677 | -36.0502 | -38.9382
23.8738 | -36.0507 | -38.9441
23.8799 | -36.0513 | -38.9502
23.886 | -36.0518 | -38.9562
23.8921 | -36.0524 | -38.9626
23.8982 | -36.053 | -38.9687
23.9043 | -36.0536 | -38.975
23.9104 | -36.0541 | -38.981
23.9165 | -36.0547 | -38.9867
23.9226 | -36.0552 | -38.9927
23.9287 | -36.0558 | -38.9988
23.9348 | -36.0564 | -39.0047
23.9409 | -36.057 | -39.0111
23.947 | -36.0576 | -39.0169
23.9531 | -36.0582 | -39.0225
23.9592 | -36.0587 | -39.0285
23.9653 | -36.0593 | -39.0347
23.9714 | -36.0598 | -39.0407
23.9775 | -36.0604 | -39.0472
23.9836 | -36.061 | -39.0534
23.9897 | -36.0616 | -39.0596
23.9958 | -36.0622 | -39.0658
24.002 | -36.0628 | -39.0719
24.0081 | -36.0634 | -39.0774
24.0142 | -36.0639 | -39.0837
24.0203 | -36.0645 | -39.0895
24.0264 | -36.0651 | -39.0955
24.0325 | -36.0657 | -39.1013
24.0386 | -36.0662 | -39.1075
24.0447 | -36.0668 | -39.1135
24.0508 | -36.0674 | -39.1197
24.0569 | -36.0679 | -39.1257
24.063 | -36.0684 | -39.1317
24.0691 | -36.069 | -39.1375
24.0752 | -36.0696 | -39.1435
24.0813 | -36.0702 | -39.1497
24.0874 | -36.0707 | -39.1559
24.0935 | -36.0712 | -39.1622
24.0996 | -36.0717 | -39.1681
24.1057 | -36.0723 | -39.1741
24.1118 | -36.0729 | -39.1803
24.1179 | -36.0734 | -39.1865
24.124 | -36.074 | -39.1922
24.1301 | -36.0746 | -39.198
24.1362 | -36.0752 | -39.2038
24.1424 | -36.0757 | -39.2099
24.1485 | -36.0763 | -39.2159
24.1546 | -36.0768 | -39.2221
24.1607 | -36.0774 | -39.2281
24.1668 | -36.078 | -39.2341
24.1729 | -36.0786 | -39.2403
24.179 | -36.0791 | -39.2466
24.1851 | -36.0798 | -39.2525
24.1912 | -36.0804 | -39.2591
24.1973 | -36.081 | -39.2655
24.2034 | -36.0815 | -39.2714
24.2095 | -36.0822 | -39.2777
24.2156 | -36.0826 | -39.284
24.2217 | -36.0832 | -39.2903
24.2278 | -36.0838 | -39.2961
24.2339 | -36.0844 | -39.3022
24.24 | -36.0849 | -39.3085
24.2461 | -36.0855 | -39.3148
24.2522 | -36.0861 | -39.3207
24.2583 | -36.0866 | -39.3268
24.2644 | -36.0872 | -39.333
24.2705 | -36.0878 | -39.3391
24.2766 | -36.0884 | -39.3453
24.2827 | -36.0889 | -39.3513
24.2889 | -36.0895 | -39.3574
24.295 | -36.0901 | -39.3638
24.3011 | -36.0906 | -39.3701
24.3072 | -36.0912 | -39.376
24.3133 | -36.0918 | -39.3818
24.3194 | -36.0923 | -39.3882
24.3255 | -36.0929 | -39.3944
24.3316 | -36.0935 | -39.4005
24.3377 | -36.0941 | -39.4064
24.3438 | -36.0946 | -39.4126
24.3499 | -36.0951 | -39.4185
24.356 | -36.0956 | -39.4245
24.3621 | -36.0962 | -39.4305
24.3682 | -36.0967 | -39.4366
24.3743 | -36.0973 | -39.443
24.3804 | -36.0978 | -39.4485
24.3865 | -36.0984 | -39.4549
24.3926 | -36.0989 | -39.4609
24.3987 | -36.0994 | -39.4669
24.4048 | -36.1 | -39.4729
24.4109 | -36.1005 | -39.4789
24.417 | -36.1011 | -39.485
24.4231 | -36.1016 | -39.4911
24.4293 | -36.1022 | -39.4969
24.4354 | -36.1027 | -39.5028
24.4415 | -36.1033 | -39.5088
24.4476 | -36.1038 | -39.5145
24.4537 | -36.1044 | -39.5204
24.4598 | -36.105 | -39.5265
24.4659 | -36.1055 | -39.5326
24.472 | -36.1061 | -39.5387
24.4781 | -36.1067 | -39.545
24.4842 | -36.1072 | -39.5506
24.4903 | -36.1077 | -39.5563
24.4964 | -36.1082 | -39.5622
24.5025 | -36.1088 | -39.5681
24.5086 | -36.1094 | -39.5743
24.5147 | -36.1099 | -39.5802
24.5208 | -36.1105 | -39.586
24.5269 | -36.111 | -39.592
24.533 | -36.1116 | -39.5979
24.5391 | -36.1121 | -39.6042
24.5452 | -36.1127 | -39.6102
24.5513 | -36.1133 | -39.616
24.5574 | -36.1139 | -39.6225
24.5635 | -36.1144 | -39.6283
24.5696 | -36.115 | -39.6343
24.5758 | -36.1155 | -39.6403
24.5819 | -36.1161 | -39.6465
24.588 | -36.1166 | -39.6522
24.5941 | -36.1172 | -39.6582
24.6002 | -36.1177 | -39.664
24.6063 | -36.1183 | -39.67
24.6124 | -36.1188 | -39.6764
24.6185 | -36.1194 | -39.6828
24.6246 | -36.1199 | -39.6886
24.6307 | -36.1205 | -39.6946
24.6368 | -36.121 | -39.7005
24.6429 | -36.1216 | -39.7065
24.649 | -36.1221 | -39.7123
24.6551 | -36.1227 | -39.718
24.6612 | -36.1233 | -39.7237
24.6673 | -36.1239 | -39.7299
24.6734 | -36.1244 | -39.7362
24.6795 | -36.1249 | -39.7423
24.6856 | -36.1255 | -39.7485
24.6917 | -36.1261 | -39.7545
24.6978 | -36.1266 | -39.7605
24.7039 | -36.1272 | -39.7665
24.71 | -36.1277 | -39.7726
24.7162 | -36.1282 | -39.7785
24.7223 | -36.1288 | -39.7844
24.7284 | -36.1294 | -39.7906
24.7345 | -36.13 | -39.7967
24.7406 | -36.1305 | -39.8028
24.7467 | -36.1311 | -39.809
24.7528 | -36.1316 | -39.8152
24.7589 | -36.1321 | -39.8211
24.765 | -36.1327 | -39.827
24.7711 | -36.1333 | -39.8329
24.7772 | -36.1338 | -39.8389
24.7833 | -36.1344 | -39.845
24.7894 | -36.1349 | -39.8512
24.7955 | -36.1355 | -39.8571
24.8016 | -36.136 | -39.8633
24.8077 | -36.1366 | -39.8695
24.8138 | -36.1371 | -39.8755
24.8199 | -36.1377 | -39.8817
24.826 | -36.1382 | -39.8879
24.8321 | -36.1388 | -39.8939
24.8382 | -36.1393 | -39.8998
24.8443 | -36.1398 | -39.9058
24.8504 | -36.1404 | -39.9117
24.8565 | -36.141 | -39.9179
24.8627 | -36.1415 | -39.9242
24.8688 | -36.1421 | -39.93
24.8749 | -36.1426 | -39.9358
24.881 | -36.1432 | -39.9419
24.8871 | -36.1437 | -39.948
24.8932 | -36.1443 | -39.9538
24.8993 | -36.1449 | -39.9597
24.9054 | -36.1454 | -39.9656
24.9115 | -36.146 | -39.9715
24.9176 | -36.1465 | -39.9772
24.9237 | -36.1471 | -39.9835
24.9298 | -36.1476 | -39.9893
24.9359 | -36.1482 | -39.9953
24.942 | -36.1488 | -40.0013
24.9481 | -36.1493 | -40.0071
24.9542 | -36.1499 | -40.0131
24.9603 | -36.1504 | -40.019
24.9664 | -36.1509 | -40.0253
24.9725 | -36.1515 | -40.0312
24.9786 | -36.1519 | -40.0369
24.9847 | -36.1525 | -40.0429
24.9908 | -36.153 | -40.0487
24.9969 | -36.1535 | -40.0548
25.0031 | -36.154 | -40.0608
25.0092 | -36.1546 | -40.0665
25.0153 | -36.1551 | -40.0722
25.0214 | -36.1556 | -40.0781
25.0275 | -36.1562 | -40.084
25.0336 | -36.1567 | -40.09
25.0397 | -36.1573 | -40.0962
25.0458 | -36.1578 | -40.1021
25.0519 | -36.1583 | -40.1082
25.058 | -36.1588 | -40.1137
25.0641 | -36.1593 | -40.1199
25.0702 | -36.1599 | -40.1259
25.0763 | -36.1604 | -40.1321
25.0824 | -36.1609 | -40.1379
25.0885 | -36.1614 | -40.1438
25.0946 | -36.162 | -40.1497
25.1007 | -36.1625 | -40.1558
25.1068 | -36.1631 | -40.1619
25.1129 | -36.1636 | -40.1679
25.119 | -36.1641 | -40.1737
25.1251 | -36.1646 | -40.1798
25.1312 | -36.1652 | -40.1858
25.1373 | -36.1657 | -40.1917
25.1435 | -36.1663 | -40.1978
25.1496 | -36.1668 | -40.2036
25.1557 | -36.1674 | -40.2095
25.1618 | -36.1679 | -40.2158
25.1679 | -36.1684 | -40.2218
25.174 | -36.1689 | -40.2275
25.1801 | -36.1695 | -40.2334
25.1862 | -36.17 | -40.2393
25.1923 | -36.1705 | -40.2452
25.1984 | -36.1711 | -40.2512
25.2045 | -36.1716 | -40.2572
25.2106 | -36.1722 | -40.263
25.2167 | -36.1727 | -40.2692
25.2228 | -36.1732 | -40.2749
25.2289 | -36.1738 | -40.2809
25.235 | -36.1743 | -40.2869
25.2411 | -36.1748 | -40.293
25.2472 | -36.1754 | -40.2988
25.2533 | -36.1759 | -40.305
25.2594 | -36.1765 | -40.3109
25.2655 | -36.177 | -40.317
25.2716 | -36.1776 | -40.323
25.2777 | -36.1781 | -40.3291
25.2838 | -36.1786 | -40.3352
25.29 | -36.1791 | -40.3412
25.2961 | -36.1797 | -40.347
25.3022 | -36.1802 | -40.3529
25.3083 | -36.1807 | -40.3588
25.3144 | -36.1812 | -40.3647
25.3205 | -36.1817 | -40.3706
25.3266 | -36.1822 | -40.3763
25.3327 | -36.1827 | -40.3819
25.3388 | -36.1832 | -40.3879
25.3449 | -36.1838 | -40.394
25.351 | -36.1843 | -40.3999
25.3571 | -36.1848 | -40.4057
25.3632 | -36.1853 | -40.4114
25.3693 | -36.1859 | -40.4173
25.3754 | -36.1864 | -40.4233
25.3815 | -36.1869 | -40.4291
25.3876 | -36.1874 | -40.435
25.3937 | -36.188 | -40.4408
25.3998 | -36.1885 | -40.4463
25.4059 | -36.189 | -40.4523
25.412 | -36.1896 | -40.4582
25.4181 | -36.1901 | -40.4641
25.4242 | -36.1906 | -40.4701
25.4304 | -36.1911 | -40.4761
25.4365 | -36.1917 | -40.4817
25.4426 | -36.1921 | -40.488
25.4487 | -36.1927 | -40.4941
25.4548 | -36.1932 | -40.5
25.4609 | -36.1937 | -40.5057
25.467 | -36.1942 | -40.5115
25.4731 | -36.1948 | -40.5175
25.4792 | -36.1953 | -40.5231
25.4853 | -36.1959 | -40.5287
25.4914 | -36.1964 | -40.5347
25.4975 | -36.1969 | -40.5406
25.5036 | -36.1975 | -40.5463
25.5097 | -36.198 | -40.5521
25.5158 | -36.1985 | -40.558
25.5219 | -36.199 | -40.5639
25.528 | -36.1996 | -40.5699
25.5341 | -36.2001 | -40.5756
25.5402 | -36.2006 | -40.5817
25.5463 | -36.2012 | -40.5878
25.5524 | -36.2017 | -40.5939
25.5585 | -36.2022 | -40.5996
25.5646 | -36.2027 | -40.6056
25.5707 | -36.2032 | -40.6117
25.5769 | -36.2037 | -40.6174
25.583 | -36.2043 | -40.6234
25.5891 | -36.2048 | -40.6294
25.5952 | -36.2054 | -40.635
25.6013 | -36.2059 | -40.6411
25.6074 | -36.2064 | -40.6468
25.6135 | -36.207 | -40.6529
25.6196 | -36.2075 | -40.6591
25.6257 | -36.208 | -40.665
25.6318 | -36.2085 | -40.6707
25.6379 | -36.209 | -40.6764
25.644 | -36.2096 | -40.6822
25.6501 | -36.2101 | -40.6879
25.6562 | -36.2106 | -40.6935
25.6623 | -36.2111 | -40.6992
25.6684 | -36.2116 | -40.7049
25.6745 | -36.2121 | -40.7108
25.6806 | -36.2127 | -40.7166
25.6867 | -36.2132 | -40.7223
25.6928 | -36.2137 | -40.7282
25.6989 | -36.2142 | -40.7338
25.705 | -36.2147 | -40.7397
25.7111 | -36.2152 | -40.7456
25.7173 | -36.2157 | -40.7512
25.7234 | -36.2162 | -40.757
25.7295 | -36.2168 | -40.7627
25.7356 | -36.2173 | -40.7687
25.7417 | -36.2178 | -40.7746
25.7478 | -36.2184 | -40.7802
25.7539 | -36.2189 | -40.7857
25.76 | -36.2195 | -40.7917
25.7661 | -36.22 | -40.7974
25.7722 | -36.2205 | -40.8031
25.7783 | -36.221 | -40.8092
25.7844 | -36.2216 | -40.8145
25.7905 | -36.222 | -40.8206
25.7966 | -36.2225 | -40.8261
25.8027 | -36.223 | -40.8318
25.8088 | -36.2235 | -40.838
25.8149 | -36.224 | -40.8439
25.821 | -36.2245 | -40.8496
25.8271 | -36.225 | -40.8554
25.8332 | -36.2255 | -40.8612
25.8393 | -36.226 | -40.8671
25.8454 | -36.2265 | -40.8728
25.8515 | -36.2271 | -40.8785
25.8576 | -36.2276 | -40.8842
25.8638 | -36.2281 | -40.89
25.8699 | -36.2287 | -40.896
25.876 | -36.2292 | -40.902
25.8821 | -36.2297 | -40.9079
25.8882 | -36.2302 | -40.9137
25.8943 | -36.2307 | -40.9198
25.9004 | -36.2312 | -40.9258
25.9065 | -36.2318 | -40.9316
25.9126 | -36.2323 | -40.9378
25.9187 | -36.2328 | -40.9438
25.9248 | -36.2333 | -40.9495
25.9309 | -36.2338 | -40.9554
25.937 | -36.2344 | -40.9612
25.9431 | -36.2349 | -40.9672
25.9492 | -36.2354 | -40.9728
25.9553 | -36.2359 | -40.9789
25.9614 | -36.2364 | -40.9845
25.9675 | -36.2369 | -40.9901
25.9736 | -36.2373 | -40.9961
25.9797 | -36.2379 | -41.0023
25.9858 | -36.2384 | -41.0078
25.9919 | -36.2389 | -41.0139
25.998 | -36.2394 | -41.0197
26.0042 | -36.2399 | -41.0257
26.0103 | -36.2404 | -41.0318
26.0164 | -36.2409 | -41.0373
26.0225 | -36.2414 | -41.0432
26.0286 | -36.242 | -41.0492
26.0347 | -36.2425 | -41.0549
26.0408 | -36.243 | -41.0607
26.0469 | -36.2435 | -41.0663
26.053 | -36.2441 | -41.0721
26.0591 | -36.2446 | -41.0779
26.0652 | -36.2452 | -41.0833
26.0713 | -36.2457 | -41.0891
26.0774 | -36.2462 | -41.0947
26.0835 | -36.2467 | -41.1007
26.0896 | -36.2472 | -41.1064
26.0957 | -36.2477 | -41.1121
26.1018 | -36.2482 | -41.1177
26.1079 | -36.2487 | -41.1233
26.114 | -36.2492 | -41.1293
26.1201 | -36.2497 | -41.1349
26.1262 | -36.2502 | -41.1407
26.1323 | -36.2507 | -41.1462
26.1384 | -36.2512 | -41.1518
26.1445 | -36.2517 | -41.1576
26.1507 | -36.2523 | -41.1634
26.1568 | -36.2528 | -41.1692
26.1629 | -36.2533 | -41.1749
26.169 | -36.2538 | -41.1805
26.1751 | -36.2544 | -41.1863
26.1812 | -36.2548 | -41.1919
26.1873 | -36.2553 | -41.1978
26.1934 | -36.2559 | -41.2032
26.1995 | -36.2564 | -41.2088
26.2056 | -36.2569 | -41.2143
26.2117 | -36.2574 | -41.2199
26.2178 | -36.2579 | -41.2257
26.2239 | -36.2585 | -41.2312
26.23 | -36.259 | -41.2373
26.2361 | -36.2595 | -41.2431
26.2422 | -36.2599 | -41.2488
26.2483 | -36.2604 | -41.2545
26.2544 | -36.261 | -41.2601
26.2605 | -36.2615 | -41.2659
26.2666 | -36.262 | -41.2716
26.2727 | -36.2624 | -41.2774
26.2788 | -36.2629 | -41.2827
26.2849 | -36.2634 | -41.2883
26.2911 | -36.2639 | -41.2938
26.2972 | -36.2644 | -41.2998
26.3033 | -36.2649 | -41.3054
26.3094 | -36.2654 | -41.3113
26.3155 | -36.2659 | -41.3167
26.3216 | -36.2664 | -41.3226
26.3277 | -36.2669 | -41.328
26.3338 | -36.2674 | -41.3335
26.3399 | -36.268 | -41.339
26.346 | -36.2685 | -41.3445
26.3521 | -36.2689 | -41.3502
26.3582 | -36.2694 | -41.3559
26.3643 | -36.27 | -41.3615
26.3704 | -36.2705 | -41.3677
26.3765 | -36.271 | -41.3736
26.3826 | -36.2715 | -41.3794
26.3887 | -36.272 | -41.3851
26.3948 | -36.2725 | -41.3906
26.4009 | -36.273 | -41.3963
26.407 | -36.2735 | -41.4021
26.4131 | -36.2739 | -41.4075
26.4192 | -36.2743 | -41.413
26.4253 | -36.2749 | -41.4183
26.4314 | -36.2754 | -41.4238
26.4376 | -36.2759 | -41.4296
26.4437 | -36.2764 | -41.4352
26.4498 | -36.2769 | -41.4406
26.4559 | -36.2774 | -41.4461
26.462 | -36.278 | -41.4516
26.4681 | -36.2784 | -41.4574
26.4742 | -36.2789 | -41.4631
26.4803 | -36.2794 | -41.4686
26.4864 | -36.2799 | -41.474
26.4925 | -36.2803 | -41.4799
26.4986 | -36.2808 | -41.4859
26.5047 | -36.2813 | -41.4916
26.5108 | -36.2818 | -41.4974
26.5169 | -36.2823 | -41.503
26.523 | -36.2827 | -41.5084
26.5291 | -36.2832 | -41.5138
26.5352 | -36.2837 | -41.5195
26.5413 | -36.2842 | -41.5253
26.5474 | -36.2847 | -41.5313
26.5535 | -36.2852 | -41.5369
26.5596 | -36.2857 | -41.5426
26.5657 | -36.2861 | -41.5482
26.5718 | -36.2866 | -41.5541
26.578 | -36.2871 | -41.56
26.5841 | -36.2876 | -41.5654
26.5902 | -36.2881 | -41.5711
26.5963 | -36.2886 | -41.5767
26.6024 | -36.2891 | -41.5823
26.6085 | -36.2895 | -41.5882
26.6146 | -36.29 | -41.5939
26.6207 | -36.2905 | -41.5996
26.6268 | -36.291 | -41.6057
26.6329 | -36.2914 | -41.6113
26.639 | -36.2919 | -41.617
26.6451 | -36.2923 | -41.6228
26.6512 | -36.2929 | -41.6281
26.6573 | -36.2934 | -41.6338
26.6634 | -36.2938 | -41.6391
26.6695 | -36.2943 | -41.645
26.6756 | -36.2948 | -41.6507
26.6817 | -36.2953 | -41.6563
26.6878 | -36.2958 | -41.6616
26.6939 | -36.2963 | -41.6673
26.7 | -36.2967 | -41.6728
26.7061 | -36.2972 | -41.6787
26.7122 | -36.2976 | -41.6845
26.7183 | -36.2981 | -41.6904
26.7245 | -36.2987 | -41.696
26.7306 | -36.2992 | -41.7019
26.7367 | -36.2997 | -41.7072
26.7428 | -36.3002 | -41.7133
26.7489 | -36.3007 | -41.719
26.755 | -36.3012 | -41.7246
26.7611 | -36.3016 | -41.7305
26.7672 | -36.3021 | -41.736
26.7733 | -36.3026 | -41.7419
26.7794 | -36.3031 | -41.7477
26.7855 | -36.3036 | -41.7534
26.7916 | -36.3041 | -41.7586
26.7977 | -36.3046 | -41.7638
26.8038 | -36.3051 | -41.7693
26.8099 | -36.3056 | -41.7754
26.816 | -36.3061 | -41.781
26.8221 | -36.3066 | -41.7865
26.8282 | -36.3071 | -41.7921
26.8343 | -36.3075 | -41.7977
26.8404 | -36.308 | -41.803
26.8465 | -36.3085 | -41.8086
26.8526 | -36.309 | -41.814
26.8587 | -36.3095 | -41.8196
26.8649 | -36.31 | -41.8252
26.871 | -36.3105 | -41.8311
26.8771 | -36.3109 | -41.8364
26.8832 | -36.3114 | -41.8423
26.8893 | -36.3119 | -41.8482
26.8954 | -36.3124 | -41.8537
26.9015 | -36.3129 | -41.8597
26.9076 | -36.3134 | -41.8656
26.9137 | -36.3138 | -41.8709
26.9198 | -36.3143 | -41.8766
26.9259 | -36.3148 | -41.882
26.932 | -36.3153 | -41.8873
26.9381 | -36.3158 | -41.8931
26.9442 | -36.3162 | -41.8983
26.9503 | -36.3167 | -41.9037
26.9564 | -36.3172 | -41.9096
26.9625 | -36.3177 | -41.9153
26.9686 | -36.3182 | -41.9208
26.9747 | -36.3187 | -41.9265
26.9808 | -36.3192 | -41.9322
26.9869 | -36.3197 | -41.9378
26.993 | -36.3201 | -41.9433
26.9991 | -36.3206 | -41.9491
27.0052 | -36.321 | -41.9552
27.0114 | -36.3215 | -41.9609
27.0175 | -36.322 | -41.9662
27.0236 | -36.3224 | -41.9719
27.0297 | -36.3229 | -41.9777
27.0358 | -36.3233 | -41.9832
27.0419 | -36.3237 | -41.9888
27.048 | -36.3242 | -41.9946
27.0541 | -36.3246 | -42.0003
27.0602 | -36.3251 | -42.0063
27.0663 | -36.3256 | -42.0117
27.0724 | -36.326 | -42.0173
27.0785 | -36.3265 | -42.0231
27.0846 | -36.327 | -42.0282
27.0907 | -36.3275 | -42.0338
27.0968 | -36.3279 | -42.0395
27.1029 | -36.3284 | -42.0449
27.109 | -36.3288 | -42.0504
27.1151 | -36.3293 | -42.0557
27.1212 | -36.3297 | -42.0611
27.1273 | -36.3302 | -42.067
27.1334 | -36.3306 | -42.0728
27.1395 | -36.3312 | -42.0787
27.1456 | -36.3316 | -42.0843
27.1518 | -36.332 | -42.0897
27.1579 | -36.3325 | -42.0952
27.164 | -36.333 | -42.1007
27.1701 | -36.3335 | -42.1065
27.1762 | -36.3339 | -42.1119
27.1823 | -36.3344 | -42.1171
27.1884 | -36.3348 | -42.1229
27.1945 | -36.3353 | -42.1284
27.2006 | -36.3357 | -42.1339
27.2067 | -36.3362 | -42.1394
27.2128 | -36.3367 | -42.1451
27.2189 | -36.3371 | -42.1507
27.225 | -36.3376 | -42.1566
27.2311 | -36.338 | -42.1617
27.2372 | -36.3385 | -42.167
27.2433 | -36.339 | -42.1728
27.2494 | -36.3395 | -42.1786
27.2555 | -36.34 | -42.1845
27.2616 | -36.3404 | -42.1899
27.2677 | -36.3409 | -42.1954
27.2738 | -36.3414 | -42.2012
27.2799 | -36.3419 | -42.2067
27.286 | -36.3424 | -42.2126
27.2921 | -36.3428 | -42.2181
27.2983 | -36.3433 | -42.2233
27.3044 | -36.3437 | -42.2288
27.3105 | -36.3442 | -42.2343
27.3166 | -36.3447 | -42.24
27.3227 | -36.3452 | -42.2456
27.3288 | -36.3457 | -42.2512
27.3349 | -36.3461 | -42.2565
27.341 | -36.3465 | -42.2619
27.3471 | -36.347 | -42.2676
27.3532 | -36.3475 | -42.2732
27.3593 | -36.348 | -42.2789
27.3654 | -36.3485 | -42.2847
27.3715 | -36.349 | -42.2899
27.3776 | -36.3495 | -42.2958
27.3837 | -36.3499 | -42.3013
27.3898 | -36.3504 | -42.3068
27.3959 | -36.3508 | -42.3123
27.402 | -36.3512 | -42.3179
27.4081 | -36.3517 | -42.3234
27.4142 | -36.3522 | -42.3291
27.4203 | -36.3527 | -42.3346
27.4264 | -36.3531 | -42.3402
27.4325 | -36.3535 | -42.3457
27.4387 | -36.354 | -42.351
27.4448 | -36.3545 | -42.3566
27.4509 | -36.3549 | -42.3625
27.457 | -36.3554 | -42.3677
27.4631 | -36.3559 | -42.3735
27.4692 | -36.3563 | -42.3787
27.4753 | -36.3567 | -42.3838
27.4814 | -36.3572 | -42.3892
27.4875 | -36.3577 | -42.3949
27.4936 | -36.3582 | -42.4003
27.4997 | -36.3586 | -42.4056
27.5058 | -36.359 | -42.411
27.5119 | -36.3594 | -42.4163
27.518 | -36.3598 | -42.4221
27.5241 | -36.3603 | -42.4279
27.5302 | -36.3607 | -42.4336
27.5363 | -36.3611 | -42.4391
27.5424 | -36.3616 | -42.4445
27.5485 | -36.362 | -42.4501
27.5546 | -36.3625 | -42.4554
27.5607 | -36.363 | -42.4611
27.5668 | -36.3634 | -42.4663
27.5729 | -36.3639 | -42.4721
27.5791 | -36.3644 | -42.4772
27.5852 | -36.3649 | -42.4824
27.5913 | -36.3654 | -42.4882
27.5974 | -36.3659 | -42.494
27.6035 | -36.3663 | -42.4993
27.6096 | -36.3667 | -42.5049
27.6157 | -36.3672 | -42.5101
27.6218 | -36.3677 | -42.5158
27.6279 | -36.3682 | -42.5212
27.634 | -36.3687 | -42.5268
27.6401 | -36.3691 | -42.5324
27.6462 | -36.3696 | -42.5379
27.6523 | -36.37 | -42.5432
27.6584 | -36.3704 | -42.5487
27.6645 | -36.3709 | -42.5538
27.6706 | -36.3714 | -42.5593
27.6767 | -36.3718 | -42.5648
27.6828 | -36.3723 | -42.5701
27.6889 | -36.3728 | -42.5756
27.695 | -36.3732 | -42.5812
27.7011 | -36.3738 | -42.5867
27.7072 | -36.3742 | -42.5925
27.7133 | -36.3747 | -42.5979
27.7194 | -36.3751 | -42.6033
27.7256 | -36.3755 | -42.6088
27.7317 | -36.3759 | -42.6141
27.7378 | -36.3763 | -42.6198
27.7439 | -36.3768 | -42.6257
27.75 | -36.3773 | -42.6315
27.7561 | -36.3778 | -42.6367
27.7622 | -36.3782 | -42.6422
27.7683 | -36.3786 | -42.6476
27.7744 | -36.3791 | -42.6529
27.7805 | -36.3795 | -42.6586
27.7866 | -36.3799 | -42.6642
27.7927 | -36.3803 | -42.6693
27.7988 | -36.3808 | -42.6747
27.8049 | -36.3813 | -42.6799
27.811 | -36.3817 | -42.6852
27.8171 | -36.3821 | -42.6904
27.8232 | -36.3826 | -42.6955
27.8293 | -36.383 | -42.7012
27.8354 | -36.3835 | -42.7067
27.8415 | -36.384 | -42.7122
27.8476 | -36.3844 | -42.7177
27.8537 | -36.3848 | -42.7232
27.8598 | -36.3852 | -42.7288
27.866 | -36.3857 | -42.7339
27.8721 | -36.3861 | -42.7394
27.8782 | -36.3865 | -42.7448
27.8843 | -36.3869 | -42.75
27.8904 | -36.3874 | -42.7552
27.8965 | -36.3878 | -42.7604
27.9026 | -36.3883 | -42.7656
27.9087 | -36.3888 | -42.7714
27.9148 | -36.3891 | -42.7761
27.9209 | -36.3895 | -42.7814
27.927 | -36.39 | -42.787
27.9331 | -36.3904 | -42.7928
27.9392 | -36.3909 | -42.7983
27.9453 | -36.3913 | -42.8039
27.9514 | -36.3917 | -42.8093
27.9575 | -36.3922 | -42.8149
27.9636 | -36.3926 | -42.8201
27.9697 | -36.393 | -42.8254
27.9758 | -36.3935 | -42.8308
27.9819 | -36.3939 | -42.8363
27.988 | -36.3943 | -42.8416
27.9941 | -36.3947 | -42.8474
28.0002 | -36.3951 | -42.8528
28.0063 | -36.3955 | -42.8583
28.0125 | -36.396 | -42.8636
28.0186 | -36.3964 | -42.8688
28.0247 | -36.3969 | -42.8738
28.0308 | -36.3974 | -42.879
28.0369 | -36.3978 | -42.884
28.043 | -36.3982 | -42.8898
28.0491 | -36.3987 | -42.8953
28.0552 | -36.3991 | -42.901
28.0613 | -36.3995 | -42.9062
28.0674 | -36.4 | -42.9116
28.0735 | -36.4004 | -42.9168

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -37.2604 | -48.0886
21.4908 | -36.1274 | -40.3209
16.4948 | -35.3592 | -32.3397
12.6602 | -34.8721 | -25.3897
9.7171 | -34.5648 | -19.7934
7.4582 | -34.3525 | -15.1895
5.7244 | -34.2015 | -11.4236
4.3936 | -34.1069 | -8.3731
3.3722 | -34.0746 | -5.9659
2.5883 | -34.0665 | -4.1252
1.9866 | -34.0539 | -2.6627
1.5248 | -34.0542 | -1.4924
1.1703 | -34.0584 | -0.70548
1 | -34.0563 | -0.31561
0.89822 | -34.0479 | -0.003599
0.68941 | -34.0397 | 0.66147
0.52919 | -34.0346 | 1.1743
0.40612 | -34.0417 | 1.8328
0.31171 | -34.0436 | 2.3159
0.23928 | -34.0278 | 3.0715
0.18358 | -34.0416 | 3.5152
0.14091 | -34.0533 | 4.4036
0.10819 | -34.0696 | 5.5716
0.083017 | -34.0923 | 7.0491
0.063705 | -34.1361 | 8.92
0.048923 | -34.2012 | 11.423
0.037479 | -34.314 | 14.6126
0.028801 | -34.4895 | 18.6598
0.022078 | -34.7956 | 23.9948
0.017023 | -35.2818 | 30.3717
0.013018 | -36.0777 | 38.0013
0.010014 | -37.1807 | 45.8932

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003673 | 12.0977
0.007346 | -29.0297
0.011019 | -36.1851
0.014692 | -40.7619
0.018365 | -42.8312
0.022039 | -45.7607
0.025712 | -46.9178
0.029385 | -48.454
0.033058 | -49.6186
0.036731 | -51.3733
0.040404 | -52.854
0.044077 | -52.6289
0.04775 | -53.723
0.051423 | -55.2267
0.055096 | -55.7727
0.05877 | -56.252
0.062443 | -57.1389
0.066116 | -56.3239
0.069789 | -56.756
0.073462 | -58.3042
0.077135 | -59.2808
0.080808 | -60.4502
0.084481 | -60.2647
0.088154 | -60.9344
0.091827 | -61.1942
0.0955 | -62.0399
0.099174 | -62.7013
0.10285 | -62.5489
0.10652 | -63.128
0.11019 | -62.299
0.11387 | -63.6538
0.11754 | -65.2149
0.12121 | -64.8173
0.12488 | -65.1903
0.12856 | -66.4506
0.13223 | -66.309
0.1359 | -65.5783
0.13958 | -65.7732
0.14325 | -67.8529
0.14692 | -68.3876
0.1506 | -67.6833
0.15427 | -67.678
0.15794 | -67.2799
0.16162 | -67.747
0.16529 | -68.2962
0.16896 | -68.2451
0.17264 | -69.2707
0.17631 | -69.3884
0.17998 | -69.1237
0.18366 | -70.0689
0.18733 | -69.6558
0.191 | -70.0267
0.19467 | -70.8908
0.19835 | -70.8706
0.20202 | -71.3594
0.20569 | -71.4373
0.20937 | -70.5686
0.21304 | -70.5575
0.21671 | -71.0293
0.22039 | -70.9958
0.22406 | -71.0908
0.22773 | -71.8193
0.2314 | -72.8463
0.23508 | -72.9738
0.23875 | -71.779
0.24242 | -71.4951
0.2461 | -71.7944
0.24977 | -73.0548
0.25344 | -73.1091
0.25712 | -73.3497
0.26079 | -73.0314
0.26446 | -72.7556
0.26814 | -73.3988
0.27181 | -73.8225
0.27548 | -73.9221
0.27915 | -74.3843
0.28283 | -73.3102
0.2865 | -73.2085
0.29017 | -74.1169
0.29385 | -75.004
0.29752 | -74.4191
0.30119 | -74.2247
0.30487 | -74.4839
0.30854 | -75.3694
0.31221 | -74.8064
0.31589 | -75.3377
0.31956 | -75.685
0.32323 | -75.1889
0.32691 | -75.5172
0.33058 | -75.234
0.33425 | -75.0088
0.33792 | -75.7532
0.3416 | -76.1961
0.34527 | -76.4005
0.34894 | -75.8748
0.35262 | -77.7933
0.35629 | -77.01
0.35996 | -76.9796
0.36364 | -76.2779
0.36731 | -75.8612
0.37098 | -76.6784
0.37466 | -77.5674
0.37833 | -77.0327
0.382 | -77.6464
0.38567 | -77.9441
0.38935 | -77.9232
0.39302 | -77.5611
0.39669 | -77.1809
0.40037 | -77.2668
0.40404 | -77.5418
0.40771 | -77.4818
0.41139 | -77.5847
0.41506 | -78.0278
0.41873 | -77.6913
0.42241 | -77.4684
0.42608 | -77.4041
0.42975 | -78.2041
0.43343 | -79.1213
0.4371 | -78.8766
0.44077 | -78.6356
0.44444 | -78.5028
0.44812 | -78.0699
0.45179 | -78.7197
0.45546 | -79.4145
0.45914 | -79.4063
0.46281 | -78.8279
0.46648 | -77.9718
0.47016 | -77.8679
0.47383 | -79.3748
0.4775 | -80.2514
0.48118 | -79.6636
0.48485 | -78.5995
0.48852 | -78.8901
0.49219 | -79.6563
0.49587 | -79.6406
0.49954 | -78.8425
0.50321 | -78.9684
0.50689 | -78.2827
0.51056 | -78.9451
0.51423 | -79.6434
0.51791 | -79.7745
0.52158 | -79.6414
0.52525 | -79.8684
0.52893 | -79.631
0.5326 | -79.2441
0.53627 | -79.0302
0.53995 | -79.3008
0.54362 | -79.7313
0.54729 | -80.1324
0.55096 | -80.2372
0.55464 | -80.4056
0.55831 | -80.7559
0.56198 | -80.8602
0.56566 | -80.2831
0.56933 | -79.8623
0.573 | -80.0412
0.57668 | -80.247
0.58035 | -80.5664
0.58402 | -80.7809
0.58769 | -80.3427
0.59137 | -80.1418
0.59504 | -80.6821
0.59871 | -81.1412
0.60239 | -81.0346
0.60606 | -81.1699
0.60973 | -81.8629
0.61341 | -82.0343
0.61708 | -81.4388
0.62075 | -81.279
0.62443 | -81.4774
0.6281 | -81.2402
0.63177 | -81.241
0.63545 | -82.0058
0.63912 | -82.8265
0.64279 | -83.3751
0.64646 | -83.1822
0.65014 | -82.3037
0.65381 | -81.9978
0.65748 | -82.1959
0.66116 | -81.9768
0.66483 | -81.7722
0.6685 | -82.3228
0.67218 | -83.2011
0.67585 | -83.6125
0.67952 | -83.4577
0.6832 | -83.2896
0.68687 | -82.9712
0.69054 | -82.7062
0.69422 | -83.0313
0.69789 | -83.4528
0.70156 | -83.0455
0.70523 | -82.5372
0.70891 | -82.6807
0.71258 | -83.1057
0.71625 | -82.9428
0.71993 | -82.5336
0.7236 | -82.556
0.72727 | -83.0056
0.73095 | -83.3067
0.73462 | -83.1139
0.73829 | -82.9655
0.74196 | -82.8401
0.74564 | -83.1297
0.74931 | -83.6787
0.75298 | -83.5339
0.75666 | -83.2111
0.76033 | -83.0999
0.764 | -83.1643
0.76768 | -83.4845
0.77135 | -84.0596
0.77502 | -84.6382
0.7787 | -84.8608
0.78237 | -84.8792
0.78604 | -84.9233
0.78971 | -84.7017
0.79339 | -84.4597
0.79706 | -84.3394
0.80073 | -84.4881
0.80441 | -84.3269
0.80808 | -83.8631
0.81175 | -83.7554
0.81543 | -84.2558
0.8191 | -84.736
0.82277 | -84.7097
0.82645 | -84.7656
0.83012 | -84.9466
0.83379 | -84.8641
0.83747 | -84.8557
0.84114 | -85.1048
0.84481 | -85.2991
0.84849 | -85.2291
0.85216 | -85.0054
0.85583 | -85.2912
0.8595 | -85.4204
0.86318 | -85.4496
0.86685 | -85.548
0.87052 | -85.4333
0.8742 | -85.7986
0.87787 | -86.0546
0.88154 | -85.7046
0.88522 | -85.7014
0.88889 | -85.87
0.89256 | -85.3435
0.89624 | -84.6774
0.89991 | -84.5734
0.90358 | -84.7851
0.90725 | -85.0188
0.91093 | -85.2854
0.9146 | -85.9207
0.91827 | -86.4327
0.92195 | -86.2624
0.92562 | -85.6414
0.92929 | -85.4256
0.93297 | -85.8232
0.93664 | -86.6068
0.94031 | -86.5131
0.94398 | -85.974
0.94766 | -86.1733
0.95133 | -86.9574
0.955 | -86.8524
0.95868 | -86.2363
0.96235 | -85.9699
0.96602 | -85.9739
0.9697 | -85.7453
0.97337 | -85.6081
0.97704 | -85.8585
0.98072 | -86.4778
0.98439 | -86.5414
0.98806 | -86.258
0.99174 | -84.8649
0.99541 | -80.9476
0.99908 | -78.5405
1.0028 | -79.5339
1.0064 | -83.2668
1.0101 | -86.7378
1.0138 | -86.9326
1.0174 | -86.2489
1.0211 | -85.898
1.0248 | -86.1843
1.0285 | -86.641
1.0321 | -86.691
1.0358 | -86.5768
1.0395 | -87.0241
1.0432 | -87.4339
1.0468 | -87.1751
1.0505 | -86.9788
1.0542 | -86.8882
1.0579 | -87.2064
1.0615 | -87.3989
1.0652 | -87.2354
1.0689 | -86.7955
1.0725 | -86.5294
1.0762 | -86.6043
1.0799 | -86.8632
1.0836 | -87.1303
1.0872 | -87.3674
1.0909 | -87.5238
1.0946 | -87.4867
1.0983 | -87.4158
1.1019 | -87.6062
1.1056 | -87.8723
1.1093 | -87.7014
1.1129 | -87.2099
1.1166 | -86.9412
1.1203 | -87.0761
1.124 | -87.5115
1.1276 | -88.011
1.1313 | -88.2633
1.135 | -88.2263
1.1387 | -88.0994
1.1423 | -87.9285
1.146 | -87.7536
1.1497 | -87.7305
1.1534 | -87.8748
1.157 | -88.0703
1.1607 | -88.2661
1.1644 | -88.4008
1.168 | -88.3283
1.1717 | -88.0713
1.1754 | -87.9027
1.1791 | -87.9832
1.1827 | -88.1585
1.1864 | -88.307
1.1901 | -88.4738
1.1938 | -88.5896
1.1974 | -88.5992
1.2011 | -88.5848
1.2048 | -88.5826
1.2084 | -88.6917
1.2121 | -88.8992
1.2158 | -88.8864
1.2195 | -88.5504
1.2231 | -88.3334
1.2268 | -88.5799
1.2305 | -89.1424
1.2342 | -89.5144
1.2378 | -89.3146
1.2415 | -88.756
1.2452 | -88.3438
1.2489 | -88.2713
1.2525 | -88.4834
1.2562 | -88.9333
1.2599 | -89.4468
1.2635 | -89.7415
1.2672 | -89.7906
1.2709 | -89.7136
1.2746 | -89.5278
1.2782 | -89.3349
1.2819 | -89.3008
1.2856 | -89.2862
1.2893 | -89.2256
1.2929 | -89.3145
1.2966 | -89.4369
1.3003 | -89.3213
1.3039 | -89.059
1.3076 | -88.8994
1.3113 | -88.9192
1.315 | -89.0369
1.3186 | -89.2428
1.3223 | -89.5226
1.326 | -89.6621
1.3297 | -89.542
1.3333 | -89.4455
1.337 | -89.6347
1.3407 | -90.0063
1.3444 | -90.2581
1.348 | -90.2501
1.3517 | -90.1173
1.3554 | -89.8976
1.359 | -89.56
1.3627 | -89.2599
1.3664 | -89.1584
1.3701 | -89.285
1.3737 | -89.5215
1.3774 | -89.5704
1.3811 | -89.3626
1.3848 | -89.2586
1.3884 | -89.4693
1.3921 | -89.8077
1.3958 | -89.974
1.3994 | -89.9285
1.4031 | -89.9239
1.4068 | -90.1338
1.4105 | -90.4398
1.4141 | -90.5652
1.4178 | -90.2953
1.4215 | -89.7585
1.4252 | -89.373
1.4288 | -89.4328
1.4325 | -89.871
1.4362 | -90.2386
1.4399 | -90.1773
1.4435 | -89.9594
1.4472 | -89.9254
1.4509 | -90.0222
1.4545 | -90.0484
1.4582 | -89.9958
1.4619 | -90.0051
1.4656 | -90.026
1.4692 | -89.8999
1.4729 | -89.7338
1.4766 | -89.7056
1.4803 | -89.7694
1.4839 | -89.8433
1.4876 | -89.9239
1.4913 | -89.9816
1.4949 | -90.072
1.4986 | -90.2617
1.5023 | -90.4065
1.506 | -90.3905
1.5096 | -90.3955
1.5133 | -90.5695
1.517 | -90.7134
1.5207 | -90.5238
1.5243 | -90.1624
1.528 | -90.0968
1.5317 | -90.4749
1.5354 | -90.9044
1.539 | -90.9753
1.5427 | -90.9049
1.5464 | -90.978
1.55 | -91.0456
1.5537 | -90.8973
1.5574 | -90.7029
1.5611 | -90.7221
1.5647 | -90.8997
1.5684 | -90.9995
1.5721 | -90.955
1.5758 | -90.7747
1.5794 | -90.5444
1.5831 | -90.5756
1.5868 | -90.9535
1.5904 | -91.1995
1.5941 | -91.332
1.5978 | -91.3595
1.6015 | -91.3803
1.6051 | -91.4497
1.6088 | -91.5321
1.6125 | -91.4977
1.6162 | -91.2983
1.6198 | -91.0498
1.6235 | -90.8902
1.6272 | -90.8439
1.6309 | -90.8274
1.6345 | -90.7749
1.6382 | -90.7393
1.6419 | -90.8007
1.6455 | -90.9789
1.6492 | -91.2088
1.6529 | -91.3498
1.6566 | -91.334
1.6602 | -91.246
1.6639 | -91.2273
1.6676 | -91.2934
1.6713 | -91.3578
1.6749 | -91.3657
1.6786 | -91.3458
1.6823 | -91.3478
1.686 | -91.3806
1.6896 | -91.4099
1.6933 | -91.406
1.697 | -91.3387
1.7006 | -91.2297
1.7043 | -91.131
1.708 | -91.0668
1.7117 | -91.0497
1.7153 | -91.1048
1.719 | -91.2303
1.7227 | -91.3477
1.7264 | -91.4142
1.73 | -91.4842
1.7337 | -91.6432
1.7374 | -91.9125
1.741 | -92.2466
1.7447 | -92.5419
1.7484 | -92.6927
1.7521 | -92.6175
1.7557 | -92.3505
1.7594 | -92.0126
1.7631 | -91.7281
1.7668 | -91.5418
1.7704 | -91.4525
1.7741 | -91.4119
1.7778 | -91.3779
1.7815 | -91.3273
1.7851 | -91.2664
1.7888 | -91.2196
1.7925 | -91.2144
1.7961 | -91.2531
1.7998 | -91.3366
1.8035 | -91.5018
1.8072 | -91.7839
1.8108 | -92.1069
1.8145 | -92.3407
1.8182 | -92.465
1.8219 | -92.5759
1.8255 | -92.7092
1.8292 | -92.7677
1.8329 | -92.6251
1.8365 | -92.2547
1.8402 | -91.7973
1.8439 | -91.4511
1.8476 | -91.3278
1.8512 | -91.4349
1.8549 | -91.6912
1.8586 | -91.9362
1.8623 | -91.9737
1.8659 | -91.784
1.8696 | -91.5693
1.8733 | -91.5241
1.877 | -91.7007
1.8806 | -92.0249
1.8843 | -92.3558
1.888 | -92.5617
1.8916 | -92.5658
1.8953 | -92.4411
1.899 | -92.3308
1.9027 | -92.3541
1.9063 | -92.5241
1.91 | -92.7787
1.9137 | -92.9961
1.9174 | -93.0679
1.921 | -93.024
1.9247 | -92.9973
1.9284 | -93.0271
1.932 | -93.078
1.9357 | -93.0939
1.9394 | -93.0714
1.9431 | -93.0092
1.9467 | -92.9473
1.9504 | -92.9383
1.9541 | -92.9911
1.9578 | -93.0887
1.9614 | -93.1914
1.9651 | -93.1921
1.9688 | -93.0326
1.9725 | -92.7994
1.9761 | -92.6233
1.9798 | -92.5367
1.9835 | -92.1386
1.9871 | -91.0816
1.9908 | -89.4407
1.9945 | -87.7927
1.9982 | -86.8003
2.0018 | -86.764
2.0055 | -87.6934
2.0092 | -89.3338
2.0129 | -91.1152
2.0165 | -92.4614
2.0202 | -93.0604
2.0239 | -93.1056
2.0275 | -92.9654
2.0312 | -92.757
2.0349 | -92.654
2.0386 | -92.7358
2.0422 | -92.9481
2.0459 | -93.1664
2.0496 | -93.2802
2.0533 | -93.2237
2.0569 | -93.0122
2.0606 | -92.7681
2.0643 | -92.6256
2.068 | -92.6217
2.0716 | -92.716
2.0753 | -92.8569
2.079 | -93.0088
2.0826 | -93.1288
2.0863 | -93.1474
2.09 | -93.0464
2.0937 | -92.8984
2.0973 | -92.7679
2.101 | -92.7228
2.1047 | -92.7769
2.1084 | -92.8892
2.112 | -92.9596
2.1157 | -92.8767
2.1194 | -92.619
2.123 | -92.4472
2.1267 | -92.3298
2.1304 | -92.3224
2.1341 | -92.4343
2.1377 | -92.6285
2.1414 | -92.8507
2.1451 | -93.0327
2.1488 | -93.1466
2.1524 | -93.1973
2.1561 | -93.1957
2.1598 | -93.1678
2.1635 | -93.1626
2.1671 | -93.2178
2.1708 | -93.3332
2.1745 | -93.507
2.1781 | -93.6764
2.1818 | -93.7718
2.1855 | -93.7592
2.1892 | -93.6264
2.1928 | -93.4261
2.1965 | -93.2519
2.2002 | -93.1899
2.2039 | -93.2545
2.2075 | -93.419
2.2112 | -93.5854
2.2149 | -93.6706
2.2185 | -93.6691
2.2222 | -93.6317
2.2259 | -93.564
2.2296 | -93.4715
2.2332 | -93.3655
2.2369 | -93.2819
2.2406 | -93.278
2.2443 | -93.3912
2.2479 | -93.586
2.2516 | -93.8073
2.2553 | -93.9733
2.259 | -94.0244
2.2626 | -93.9537
2.2663 | -93.827
2.27 | -93.7194
2.2736 | -93.6882
2.2773 | -93.7637
2.281 | -93.9463
2.2847 | -94.1989
2.2883 | -94.4601
2.292 | -94.6623
2.2957 | -94.7787
2.2994 | -94.8467
2.303 | -94.9007
2.3067 | -94.9592
2.3104 | -94.9943
2.314 | -94.9612
2.3177 | -94.7959
2.3214 | -94.4811
2.3251 | -94.1163
2.3287 | -93.8459
2.3324 | -93.7813
2.3361 | -93.9358
2.3398 | -94.2341
2.3434 | -94.569
2.3471 | -94.8516
2.3508 | -95.0408
2.3545 | -95.1171
2.3581 | -95.0945
2.3618 | -95.0184
2.3655 | -94.9183
2.3691 | -94.8165
2.3728 | -94.7044
2.3765 | -94.5668
2.3802 | -94.3911
2.3838 | -94.172
2.3875 | -93.9464
2.3912 | -93.7724
2.3949 | -93.6988
2.3985 | -93.7389
2.4022 | -93.852
2.4059 | -94.0004
2.4095 | -94.1359
2.4132 | -94.1936
2.4169 | -94.1571
2.4206 | -94.0615
2.4242 | -93.9784
2.4279 | -93.9583
2.4316 | -93.9874
2.4353 | -94.0261
2.4389 | -94.0529
2.4426 | -94.0566
2.4463 | -94.0385
2.45 | -94.0397
2.4536 | -94.1439
2.4573 | -94.3534
2.461 | -94.6104
2.4646 | -94.8212
2.4683 | -94.9267
2.472 | -94.9705
2.4757 | -94.9948
2.4793 | -94.9929
2.483 | -94.9535
2.4867 | -94.8981
2.4904 | -94.8498
2.494 | -94.809
2.4977 | -94.7684
2.5014 | -94.7079
2.5051 | -94.6347
2.5087 | -94.5733
2.5124 | -94.557
2.5161 | -94.6059
2.5197 | -94.7178
2.5234 | -94.8444
2.5271 | -94.9265
2.5308 | -94.9475
2.5344 | -94.8946
2.5381 | -94.7826
2.5418 | -94.6394
2.5455 | -94.5118
2.5491 | -94.4476
2.5528 | -94.436
2.5565 | -94.4569
2.5601 | -94.4894
2.5638 | -94.521
2.5675 | -94.5573
2.5712 | -94.5845
2.5748 | -94.6134
2.5785 | -94.645
2.5822 | -94.7003
2.5859 | -94.7973
2.5895 | -94.9014
2.5932 | -94.9724
2.5969 | -94.9857
2.6006 | -94.9175
2.6042 | -94.7885
2.6079 | -94.6289
2.6116 | -94.5
2.6152 | -94.4484
2.6189 | -94.4862
2.6226 | -94.5799
2.6263 | -94.6921
2.6299 | -94.8087
2.6336 | -94.9282
2.6373 | -95.0367
2.641 | -95.112
2.6446 | -95.0989
2.6483 | -94.9389
2.652 | -94.8149
2.6556 | -94.6765
2.6593 | -94.5404
2.663 | -94.4419
2.6667 | -94.3864
2.6703 | -94.3779
2.674 | -94.4307
2.6777 | -94.5372
2.6814 | -94.674
2.685 | -94.8158
2.6887 | -94.9512
2.6924 | -95.0675
2.6961 | -95.1438
2.6997 | -95.1829
2.7034 | -95.197
2.7071 | -95.2003
2.7107 | -95.2132
2.7144 | -95.2614
2.7181 | -95.3336
2.7218 | -95.3942
2.7254 | -95.4148
2.7291 | -95.3888
2.7328 | -95.3294
2.7365 | -95.2603
2.7401 | -95.2056
2.7438 | -95.1733
2.7475 | -95.1478
2.7511 | -95.1062
2.7548 | -95.0541
2.7585 | -95.0077
2.7622 | -94.9675
2.7658 | -94.9417
2.7695 | -94.9256
2.7732 | -94.9253
2.7769 | -94.9404
2.7805 | -94.9723
2.7842 | -95.0276
2.7879 | -95.0905
2.7916 | -95.1406
2.7952 | -95.1651
2.7989 | -95.1659
2.8026 | -95.1573
2.8062 | -95.159
2.8099 | -95.1938
2.8136 | -95.2473
2.8173 | -95.2992
2.8209 | -95.331
2.8246 | -95.3724
2.8283 | -95.432
2.832 | -95.5151
2.8356 | -95.6109
2.8393 | -95.6698
2.843 | -95.6471
2.8466 | -95.5331
2.8503 | -95.3688
2.854 | -95.2151
2.8577 | -95.1134
2.8613 | -95.0898
2.865 | -95.1399
2.8687 | -95.2396
2.8724 | -95.3605
2.876 | -95.4843
2.8797 | -95.5878
2.8834 | -95.6717
2.8871 | -95.7409
2.8907 | -95.7974
2.8944 | -95.8527
2.8981 | -95.9042
2.9017 | -95.9289
2.9054 | -95.9124
2.9091 | -95.8605
2.9128 | -95.8039
2.9164 | -95.7697
2.9201 | -95.7734
2.9238 | -95.8206
2.9275 | -95.8826
2.9311 | -95.9413
2.9348 | -95.9916
2.9385 | -96.023
2.9421 | -96.0206
2.9458 | -95.9946
2.9495 | -95.9658
2.9532 | -95.9419
2.9568 | -95.9206
2.9605 | -95.8973
2.9642 | -95.8737
2.9679 | -95.8448
2.9715 | -95.8084
2.9752 | -95.7344
2.9789 | -95.5997
2.9826 | -95.4171
2.9862 | -95.1822
2.9899 | -94.9186
2.9936 | -94.6686
2.9972 | -94.4983
3.0009 | -94.454
3.0046 | -94.5542
3.0083 | -94.7759
3.0119 | -95.0485
3.0156 | -95.3168
3.0193 | -95.5234
3.023 | -95.6495
3.0266 | -95.6917
3.0303 | -95.6741
3.034 | -95.672
3.0376 | -95.7185
3.0413 | -95.8124
3.045 | -95.9272
3.0487 | -96.0191
3.0523 | -96.0652
3.056 | -96.0636
3.0597 | -96.0475
3.0634 | -96.0339
3.067 | -96.061
3.0707 | -96.1321
3.0744 | -96.2423
3.0781 | -96.3743
3.0817 | -96.4979
3.0854 | -96.5978
3.0891 | -96.6702
3.0927 | -96.7151
3.0964 | -96.7482
3.1001 | -96.7726
3.1038 | -96.7953
3.1074 | -96.8131
3.1111 | -96.8221
3.1148 | -96.8377
3.1185 | -96.8708
3.1221 | -96.8885
3.1258 | -96.877
3.1295 | -96.8272
3.1332 | -96.7319
3.1368 | -96.6149
3.1405 | -96.4875
3.1442 | -96.3621
3.1478 | -96.2357
3.1515 | -96.1199
3.1552 | -96.0522
3.1589 | -96.0534
3.1625 | -96.1236
3.1662 | -96.2563
3.1699 | -96.4322
3.1736 | -96.6117
3.1772 | -96.665
3.1809 | -96.7125
3.1846 | -96.6984
3.1882 | -96.6282
3.1919 | -96.5304
3.1956 | -96.4198
3.1993 | -96.3303
3.2029 | -96.2751
3.2066 | -96.2265
3.2103 | -96.1883
3.214 | -96.1478
3.2176 | -96.0987
3.2213 | -96.0458
3.225 | -95.9953
3.2287 | -95.9719
3.2323 | -95.9838
3.236 | -96.0462
3.2397 | -96.162
3.2433 | -96.3286
3.247 | -96.5349
3.2507 | -96.7619
3.2544 | -96.9744
3.258 | -97.1383
3.2617 | -97.2452
3.2654 | -97.2738
3.2691 | -97.2488
3.2727 | -97.1964
3.2764 | -97.1342
3.2801 | -97.0807
3.2837 | -97.033
3.2874 | -96.9932
3.2911 | -96.9576
3.2948 | -96.9168
3.2984 | -96.8683
3.3021 | -96.8249
3.3058 | -96.7913
3.3095 | -96.7765
3.3131 | -96.7881
3.3168 | -96.8149
3.3205 | -96.8467
3.3242 | -96.8688
3.3278 | -96.8724
3.3315 | -96.855
3.3352 | -96.8117
3.3388 | -96.757
3.3425 | -96.71
3.3462 | -96.6831
3.3499 | -96.6923
3.3535 | -96.7388
3.3572 | -96.807
3.3609 | -96.8846
3.3646 | -96.9502
3.3682 | -96.9946
3.3719 | -97.0196
3.3756 | -97.0271
3.3792 | -97.0232
3.3829 | -97.0254
3.3866 | -97.0397
3.3903 | -97.0692
3.3939 | -97.1177
3.3976 | -97.1925
3.4013 | -97.2941
3.405 | -97.4025
3.4086 | -97.4995
3.4123 | -97.583
3.416 | -97.647
3.4197 | -97.688
3.4233 | -97.7105
3.427 | -97.73
3.4307 | -97.7453
3.4343 | -97.7618
3.438 | -97.7768
3.4417 | -97.8085
3.4454 | -97.8668
3.449 | -97.9497
3.4527 | -98.05
3.4564 | -98.1605
3.4601 | -98.2489
3.4637 | -98.2906
3.4674 | -98.2551
3.4711 | -98.1414
3.4747 | -97.9747
3.4784 | -97.7849
3.4821 | -97.5924
3.4858 | -97.4194
3.4894 | -97.277
3.4931 | -97.1747
3.4968 | -97.114
3.5005 | -97.0974
3.5041 | -97.1282
3.5078 | -97.1906
3.5115 | -97.2627
3.5152 | -97.3337
3.5188 | -97.3863
3.5225 | -97.4001
3.5262 | -97.3713
3.5298 | -97.2914
3.5335 | -97.1829
3.5372 | -97.0705
3.5409 | -96.9797
3.5445 | -96.9333
3.5482 | -96.9498
3.5519 | -97.0208
3.5556 | -97.135
3.5592 | -97.2712
3.5629 | -97.4038
3.5666 | -97.5044
3.5702 | -97.5525
3.5739 | -97.547
3.5776 | -97.5232
3.5813 | -97.4896
3.5849 | -97.4691
3.5886 | -97.4599
3.5923 | -97.4692
3.596 | -97.4934
3.5996 | -97.5189
3.6033 | -97.5321
3.607 | -97.535
3.6107 | -97.5249
3.6143 | -97.4995
3.618 | -97.4667
3.6217 | -97.4313
3.6253 | -97.4016
3.629 | -97.3902
3.6327 | -97.3984
3.6364 | -97.4275
3.64 | -97.4753
3.6437 | -97.5462
3.6474 | -97.6253
3.6511 | -97.6941
3.6547 | -97.7643
3.6584 | -97.8292
3.6621 | -97.8657
3.6657 | -97.8665
3.6694 | -97.8458
3.6731 | -97.8097
3.6768 | -97.7544
3.6804 | -97.7016
3.6841 | -97.6735
3.6878 | -97.6671
3.6915 | -97.6786
3.6951 | -97.7008
3.6988 | -97.712
3.7025 | -97.716
3.7062 | -97.7117
3.7098 | -97.7212
3.7135 | -97.7323
3.7172 | -97.7484
3.7208 | -97.7726
3.7245 | -97.7981
3.7282 | -97.8051
3.7319 | -97.7978
3.7355 | -97.7797
3.7392 | -97.7333
3.7429 | -97.6672
3.7466 | -97.5994
3.7502 | -97.5475
3.7539 | -97.504
3.7576 | -97.48
3.7612 | -97.4757
3.7649 | -97.4836
3.7686 | -97.5006
3.7723 | -97.5066
3.7759 | -97.4917
3.7796 | -97.4614
3.7833 | -97.4115
3.787 | -97.3512
3.7906 | -97.2967
3.7943 | -97.2563
3.798 | -97.2516
3.8017 | -97.2853
3.8053 | -97.3561
3.809 | -97.4559
3.8127 | -97.5741
3.8163 | -97.6933
3.82 | -97.7976
3.8237 | -97.8668
3.8274 | -97.8914
3.831 | -97.8816
3.8347 | -97.8559
3.8384 | -97.837
3.8421 | -97.8334
3.8457 | -97.8605
3.8494 | -97.9258
3.8531 | -98.0209
3.8567 | -98.1285
3.8604 | -98.2264
3.8641 | -98.3096
3.8678 | -98.3546
3.8714 | -98.3588
3.8751 | -98.3161
3.8788 | -98.2483
3.8825 | -98.1652
3.8861 | -98.0905
3.8898 | -98.0463
3.8935 | -98.039
3.8972 | -98.0658
3.9008 | -98.1232
3.9045 | -98.2094
3.9082 | -98.2942
3.9118 | -98.356
3.9155 | -98.4063
3.9192 | -98.4373
3.9229 | -98.4552
3.9265 | -98.4681
3.9302 | -98.4837
3.9339 | -98.5027
3.9376 | -98.524
3.9412 | -98.5448
3.9449 | -98.5709
3.9486 | -98.5957
3.9522 | -98.624
3.9559 | -98.6593
3.9596 | -98.6963
3.9633 | -98.7139
3.9669 | -98.6573
3.9706 | -98.4016
3.9743 | -98.0958
3.978 | -97.6927
3.9816 | -97.202
3.9853 | -96.6651
3.989 | -96.1518
3.9927 | -95.7281
3.9963 | -95.4472
4 | -95.3454
4.0037 | -95.4379
4.0073 | -95.716
4.011 | -96.1441
4.0147 | -96.6671
4.0184 | -97.211
4.022 | -97.7073
4.0257 | -98.1067
4.0294 | -98.409
4.0331 | -98.6736
4.0367 | -98.757
4.0404 | -98.7811
4.0441 | -98.8
4.0477 | -98.8032
4.0514 | -98.7846
4.0551 | -98.746
4.0588 | -98.6971
4.0624 | -98.6517
4.0661 | -98.6264
4.0698 | -98.6214
4.0735 | -98.6335
4.0771 | -98.6477
4.0808 | -98.6604
4.0845 | -98.667
4.0882 | -98.6701
4.0918 | -98.674
4.0955 | -98.6885
4.0992 | -98.7159
4.1028 | -98.7663
4.1065 | -98.8211
4.1102 | -98.8755
4.1139 | -98.9176
4.1175 | -98.9443
4.1212 | -98.9484
4.1249 | -98.9299
4.1286 | -98.8891
4.1322 | -98.836
4.1359 | -98.7865
4.1396 | -98.7464
4.1433 | -98.7275
4.1469 | -98.7421
4.1506 | -98.7939
4.1543 | -98.8821
4.1579 | -98.9783
4.1616 | -99.0682
4.1653 | -99.1306
4.169 | -99.1584
4.1726 | -99.1423
4.1763 | -99.0907
4.18 | -99.0059
4.1837 | -98.9
4.1873 | -98.7976
4.191 | -98.7011
4.1947 | -98.6268
4.1983 | -98.5725
4.202 | -98.5316
4.2057 | -98.4975
4.2094 | -98.4733
4.213 | -98.4491
4.2167 | -98.4184
4.2204 | -98.3958
4.2241 | -98.3762
4.2277 | -98.3685
4.2314 | -98.3803
4.2351 | -98.431
4.2388 | -98.5454
4.2424 | -98.6368
4.2461 | -98.7401
4.2498 | -98.8399
4.2534 | -98.9092
4.2571 | -98.9518
4.2608 | -98.9537
4.2645 | -98.9098
4.2681 | -98.8368
4.2718 | -98.7377
4.2755 | -98.6128
4.2792 | -98.4958
4.2828 | -98.3984
4.2865 | -98.3328
4.2902 | -98.3168
4.2938 | -98.3424
4.2975 | -98.4043
4.3012 | -98.4973
4.3049 | -98.6176
4.3085 | -98.7512
4.3122 | -98.8838
4.3159 | -99.0039
4.3196 | -99.1028
4.3232 | -99.178
4.3269 | -99.2249
4.3306 | -99.2488
4.3343 | -99.2628
4.3379 | -99.2718
4.3416 | -99.2908
4.3453 | -99.3243
4.3489 | -99.3602
4.3526 | -99.3881
4.3563 | -99.4064
4.36 | -99.4133
4.3636 | -99.413
4.3673 | -99.4043
4.371 | -99.3895
4.3747 | -99.3796
4.3783 | -99.3744
4.382 | -99.3815
4.3857 | -99.4087
4.3893 | -99.4591
4.393 | -99.5189
4.3967 | -99.5864
4.4004 | -99.6505
4.404 | -99.7044
4.4077 | -99.7372
4.4114 | -99.7442
4.4151 | -99.7355
4.4187 | -99.714
4.4224 | -99.6856
4.4261 | -99.6517
4.4298 | -99.6197
4.4334 | -99.5948
4.4371 | -99.5704
4.4408 | -99.5506
4.4444 | -99.5332
4.4481 | -99.5218
4.4518 | -99.513
4.4555 | -99.4974
4.4591 | -99.4712
4.4628 | -99.4369
4.4665 | -99.3881
4.4702 | -99.3347
4.4738 | -99.2784
4.4775 | -99.2243
4.4812 | -99.1888
4.4848 | -99.172
4.4885 | -99.1903
4.4922 | -99.251
4.4959 | -99.3399
4.4995 | -99.4485
4.5032 | -99.5641
4.5069 | -99.6807
4.5106 | -99.7715
4.5142 | -99.8277
4.5179 | -99.8463
4.5216 | -99.8316
4.5253 | -99.7938
4.5289 | -99.7502
4.5326 | -99.7217
4.5363 | -99.714
4.5399 | -99.7316
4.5436 | -99.7741
4.5473 | -99.8367
4.551 | -99.8971
4.5546 | -99.9487
4.5583 | -99.9785
4.562 | -99.9901
4.5657 | -99.9763
4.5693 | -99.9479
4.573 | -99.9172
4.5767 | -99.8732
4.5803 | -99.8189
4.584 | -99.7675
4.5877 | -99.7215
4.5914 | -99.6773
4.595 | -99.6442
4.5987 | -99.6332
4.6024 | -99.6379
4.6061 | -99.6616
4.6097 | -99.7013
4.6134 | -99.7534
4.6171 | -99.8112
4.6208 | -99.8706
4.6244 | -99.9251
4.6281 | -99.966
4.6318 | -99.9903
4.6354 | -100.0059
4.6391 | -100.0075
4.6428 | -99.9958
4.6465 | -99.9781
4.6501 | -99.9648
4.6538 | -99.9337
4.6575 | -99.8933
4.6612 | -99.8505
4.6648 | -99.8026
4.6685 | -99.7493
4.6722 | -99.7096
4.6758 | -99.685
4.6795 | -99.6839
4.6832 | -99.7084
4.6869 | -99.747
4.6905 | -99.7982
4.6942 | -99.8567
4.6979 | -99.9113
4.7016 | -99.9579
4.7052 | -99.9909
4.7089 | -99.9993
4.7126 | -99.994
4.7163 | -99.9817
4.7199 | -99.9667
4.7236 | -99.9641
4.7273 | -99.9616
4.7309 | -99.9636
4.7346 | -99.9699
4.7383 | -99.9727
4.742 | -99.9616
4.7456 | -99.9404
4.7493 | -99.9041
4.753 | -99.8484
4.7567 | -99.7899
4.7603 | -99.7374
4.764 | -99.687
4.7677 | -99.6656
4.7713 | -99.6499
4.775 | -99.6334
4.7787 | -99.6149
4.7824 | -99.599
4.786 | -99.5715
4.7897 | -99.5379
4.7934 | -99.5006
4.7971 | -99.4581
4.8007 | -99.4229
4.8044 | -99.3944
4.8081 | -99.37
4.8118 | -99.3501
4.8154 | -99.3323
4.8191 | -99.3161
4.8228 | -99.3058
4.8264 | -99.3109
4.8301 | -99.3324
4.8338 | -99.3694
4.8375 | -99.4317
4.8411 | -99.5105
4.8448 | -99.5771
4.8485 | -99.6406
4.8522 | -99.7006
4.8558 | -99.7401
4.8595 | -99.7521
4.8632 | -99.737
4.8669 | -99.6963
4.8705 | -99.6354
4.8742 | -99.5713
4.8779 | -99.5046
4.8815 | -99.4426
4.8852 | -99.386
4.8889 | -99.3439
4.8926 | -99.3238
4.8962 | -99.318
4.8999 | -99.3302
4.9036 | -99.3577
4.9073 | -99.4054
4.9109 | -99.469
4.9146 | -99.5402
4.9183 | -99.62
4.9219 | -99.7167
4.9256 | -99.8147
4.9293 | -99.9058
4.933 | -99.9978
4.9366 | -100.0877
4.9403 | -100.1674
4.944 | -100.23
4.9477 | -100.273
4.9513 | -100.3028
4.955 | -100.3175
4.9587 | -100.3054
4.9624 | -100.1537
4.966 | -99.9717
4.9697 | -99.7618
4.9734 | -99.4916
4.977 | -99.1715
4.9807 | -98.8194
4.9844 | -98.4595
4.9881 | -98.1299
4.9917 | -97.8621
4.9954 | -97.6746
4.9991 | -97.5831
5.0028 | -97.6051
5.0064 | -97.737
5.0101 | -97.9589
5.0138 | -98.2518
5.0174 | -98.5993
5.0211 | -98.9596
5.0248 | -99.2837
5.0285 | -99.5462
5.0321 | -99.7336
5.0358 | -99.8479
5.0395 | -99.9608
5.0432 | -99.9551
5.0468 | -99.8855
5.0505 | -99.8249
5.0542 | -99.7817
5.0579 | -99.7612
5.0615 | -99.7677
5.0652 | -99.8011
5.0689 | -99.8615
5.0725 | -99.9423
5.0762 | -100.0434
5.0799 | -100.1399
5.0836 | -100.2319
5.0872 | -100.3164
5.0909 | -100.3732
5.0946 | -100.3951
5.0983 | -100.4035
5.1019 | -100.3981
5.1056 | -100.3669
5.1093 | -100.3249
5.1129 | -100.2822
5.1166 | -100.2351
5.1203 | -100.188
5.124 | -100.1474
5.1276 | -100.1145
5.1313 | -100.0729
5.135 | -100.0362
5.1387 | -99.9982
5.1423 | -99.9655
5.146 | -99.9355
5.1497 | -99.9086
5.1534 | -99.8807
5.157 | -99.8608
5.1607 | -99.8443
5.1644 | -99.8331
5.168 | -99.8443
5.1717 | -99.8693
5.1754 | -99.9015
5.1791 | -99.945
5.1827 | -99.9995
5.1864 | -100.0544
5.1901 | -100.1069
5.1938 | -100.1606
5.1974 | -100.2104
5.2011 | -100.2602
5.2048 | -100.308
5.2084 | -100.3548
5.2121 | -100.3924
5.2158 | -100.4266
5.2195 | -100.4511
5.2231 | -100.4711
5.2268 | -100.4855
5.2305 | -100.4906
5.2342 | -100.4918
5.2378 | -100.4836
5.2415 | -100.4559
5.2452 | -100.4095
5.2489 | -100.3537
5.2525 | -100.2975
5.2562 | -100.2427
5.2599 | -100.1856
5.2635 | -100.1363
5.2672 | -100.1027
5.2709 | -100.0883
5.2746 | -100.0915
5.2782 | -100.1123
5.2819 | -100.1509
5.2856 | -100.2148
5.2893 | -100.2892
5.2929 | -100.368
5.2966 | -100.4451
5.3003 | -100.5079
5.3039 | -100.5673
5.3076 | -100.615
5.3113 | -100.6443
5.315 | -100.6537
5.3186 | -100.6449
5.3223 | -100.6203
5.326 | -100.5853
5.3297 | -100.5507
5.3333 | -100.5118
5.337 | -100.4774
5.3407 | -100.4486
5.3444 | -100.421
5.348 | -100.395
5.3517 | -100.3686
5.3554 | -100.3444
5.359 | -100.3142
5.3627 | -100.2821
5.3664 | -100.246
5.3701 | -100.2052
5.3737 | -100.1669
5.3774 | -100.1247
5.3811 | -100.088
5.3848 | -100.0568
5.3884 | -100.0242
5.3921 | -100.0079
5.3958 | -100.005
5.3994 | -100.0137
5.4031 | -100.0379
5.4068 | -100.0696
5.4105 | -100.1081
5.4141 | -100.1485
5.4178 | -100.1972
5.4215 | -100.2531
5.4252 | -100.3127
5.4288 | -100.3708
5.4325 | -100.4264
5.4362 | -100.4852
5.4399 | -100.5382
5.4435 | -100.5849
5.4472 | -100.6207
5.4509 | -100.6503
5.4545 | -100.6747
5.4582 | -100.6877
5.4619 | -100.7049
5.4656 | -100.7236
5.4692 | -100.7484
5.4729 | -100.7778
5.4766 | -100.8006
5.4803 | -100.8225
5.4839 | -100.8502
5.4876 | -100.8729
5.4913 | -100.896
5.4949 | -100.9164
5.4986 | -100.936
5.5023 | -100.9456
5.506 | -100.9492
5.5096 | -100.9497
5.5133 | -100.9509
5.517 | -100.9466
5.5207 | -100.9388
5.5243 | -100.9385
5.528 | -100.9473
5.5317 | -100.9645
5.5354 | -100.9853
5.539 | -101.0178
5.5427 | -101.0608
5.5464 | -101.1016
5.55 | -101.1457
5.5537 | -101.1817
5.5574 | -101.1971
5.5611 | -101.1938
5.5647 | -101.1666
5.5684 | -101.1141
5.5721 | -101.0558
5.5758 | -100.9919
5.5794 | -100.9257
5.5831 | -100.8625
5.5868 | -100.8086
5.5904 | -100.7643
5.5941 | -100.7308
5.5978 | -100.7111
5.6015 | -100.6997
5.6051 | -100.6924
5.6088 | -100.6916
5.6125 | -100.6929
5.6162 | -100.6982
5.6198 | -100.7058
5.6235 | -100.7157
5.6272 | -100.7224
5.6309 | -100.7284
5.6345 | -100.7371
5.6382 | -100.7476
5.6419 | -100.7679
5.6455 | -100.7884
5.6492 | -100.8091
5.6529 | -100.8236
5.6566 | -100.8364
5.6602 | -100.8556
5.6639 | -100.8752
5.6676 | -100.8959
5.6713 | -100.9145
5.6749 | -100.9291
5.6786 | -100.9401
5.6823 | -100.9486
5.686 | -100.9654
5.6896 | -100.989
5.6933 | -101.0118
5.697 | -101.043
5.7006 | -101.0785
5.7043 | -101.1163
5.708 | -101.1563
5.7117 | -101.1985
5.7153 | -101.2419
5.719 | -101.2731
5.7227 | -101.298
5.7264 | -101.3186
5.73 | -101.3318
5.7337 | -101.3355
5.7374 | -101.3365
5.741 | -101.3153
5.7447 | -101.278
5.7484 | -101.2403
5.7521 | -101.2048
5.7557 | -101.1677
5.7594 | -101.128
5.7631 | -101.0925
5.7668 | -101.0653
5.7704 | -101.0382
5.7741 | -101.0152
5.7778 | -101.0016
5.7815 | -100.9767
5.7851 | -100.9436
5.7888 | -100.9068
5.7925 | -100.866
5.7961 | -100.8238
5.7998 | -100.7822
5.8035 | -100.7436
5.8072 | -100.7068
5.8108 | -100.67
5.8145 | -100.6501
5.8182 | -100.6494
5.8219 | -100.6632
5.8255 | -100.7193
5.8292 | -100.7549
5.8329 | -100.7942
5.8365 | -100.8321
5.8402 | -100.8613
5.8439 | -100.8801
5.8476 | -100.8887
5.8512 | -100.8845
5.8549 | -100.867
5.8586 | -100.8447
5.8623 | -100.8183
5.8659 | -100.787
5.8696 | -100.7625
5.8733 | -100.7519
5.877 | -100.7487
5.8806 | -100.7543
5.8843 | -100.7721
5.888 | -100.7982
5.8916 | -100.8282
5.8953 | -100.8674
5.899 | -100.9147
5.9027 | -100.9585
5.9063 | -100.9937
5.91 | -101.0224
5.9137 | -101.0517
5.9174 | -101.0773
5.921 | -101.0956
5.9247 | -101.1129
5.9284 | -101.1277
5.932 | -101.1356
5.9357 | -101.1424
5.9394 | -101.1395
5.9431 | -101.1288
5.9467 | -101.1115
5.9504 | -101.0884
5.9541 | -101.0243
5.9578 | -100.9385
5.9614 | -100.8623
5.9651 | -100.7586
5.9688 | -100.6365
5.9725 | -100.4999
5.9761 | -100.3571
5.9798 | -100.2176
5.9835 | -100.0917
5.9871 | -99.985
5.9908 | -99.9005
5.9945 | -99.8551
5.9982 | -99.8551
6.0018 | -99.8999
6.0055 | -99.9817
6.0092 | -100.1006
6.0129 | -100.2503
6.0165 | -100.4155
6.0202 | -100.5826
6.0239 | -100.7488
6.0275 | -100.8977
6.0312 | -101.0265
6.0349 | -101.1294
6.0386 | -101.2027
6.0422 | -101.2483
6.0459 | -101.2977
6.0496 | -101.3226
6.0533 | -101.2967
6.0569 | -101.2716
6.0606 | -101.2496
6.0643 | -101.2259
6.068 | -101.2078
6.0716 | -101.1941
6.0753 | -101.1805
6.079 | -101.1567
6.0826 | -101.1476
6.0863 | -101.1409
6.09 | -101.1294
6.0937 | -101.1174
6.0973 | -101.1075
6.101 | -101.0959
6.1047 | -101.0877
6.1084 | -101.0825
6.112 | -101.081
6.1157 | -101.0875
6.1194 | -101.0967
6.123 | -101.1149
6.1267 | -101.1372
6.1304 | -101.1676
6.1341 | -101.2094
6.1377 | -101.2497
6.1414 | -101.2949
6.1451 | -101.3507
6.1488 | -101.4116
6.1524 | -101.4762
6.1561 | -101.5455
6.1598 | -101.6143
6.1635 | -101.6879
6.1671 | -101.7619
6.1708 | -101.8385
6.1745 | -101.894
6.1781 | -101.9373
6.1818 | -101.9678
6.1855 | -101.9806
6.1892 | -101.978
6.1928 | -101.9645
6.1965 | -101.9268
6.2002 | -101.8792
6.2039 | -101.8296
6.2075 | -101.7695
6.2112 | -101.7003
6.2149 | -101.6282
6.2185 | -101.5586
6.2222 | -101.4923
6.2259 | -101.4227
6.2296 | -101.3598
6.2332 | -101.2978
6.2369 | -101.2437
6.2406 | -101.2
6.2443 | -101.1659
6.2479 | -101.1458
6.2516 | -101.1318
6.2553 | -101.1282
6.259 | -101.1386
6.2626 | -101.1557
6.2663 | -101.1844
6.27 | -101.221
6.2736 | -101.266
6.2773 | -101.3116
6.281 | -101.3591
6.2847 | -101.4021
6.2883 | -101.44
6.292 | -101.4671
6.2957 | -101.4859
6.2994 | -101.4968
6.303 | -101.5098
6.3067 | -101.5153
6.3104 | -101.5193
6.314 | -101.5285
6.3177 | -101.5445
6.3214 | -101.564
6.3251 | -101.5861
6.3287 | -101.6111
6.3324 | -101.634
6.3361 | -101.6637
6.3398 | -101.6963
6.3434 | -101.7286
6.3471 | -101.7446
6.3508 | -101.7547
6.3545 | -101.7564
6.3581 | -101.7352
6.3618 | -101.7257
6.3655 | -101.7179
6.3691 | -101.7127
6.3728 | -101.7045
6.3765 | -101.694
6.3802 | -101.6929
6.3838 | -101.6949
6.3875 | -101.6977
6.3912 | -101.6919
6.3949 | -101.6855
6.3985 | -101.6781
6.4022 | -101.6663
6.4059 | -101.6498
6.4096 | -101.6303
6.4132 | -101.6145
6.4169 | -101.6022
6.4206 | -101.5886
6.4242 | -101.5749
6.4279 | -101.5676
6.4316 | -101.5566
6.4353 | -101.5489
6.4389 | -101.5464
6.4426 | -101.5442
6.4463 | -101.5461
6.45 | -101.5519
6.4536 | -101.5574
6.4573 | -101.5608
6.461 | -101.5625
6.4646 | -101.5636
6.4683 | -101.5672
6.472 | -101.5792
6.4757 | -101.5852
6.4793 | -101.591
6.483 | -101.6087
6.4867 | -101.63
6.4904 | -101.6603
6.494 | -101.6959
6.4977 | -101.7307
6.5014 | -101.7687
6.5051 | -101.8071
6.5087 | -101.8485
6.5124 | -101.8935
6.5161 | -101.9367
6.5197 | -101.9774
6.5234 | -102.0117
6.5271 | -102.0409
6.5308 | -102.0684
6.5344 | -102.0874
6.5381 | -102.0997
6.5418 | -102.1067
6.5455 | -102.091
6.5491 | -102.0611
6.5528 | -102.0262
6.5565 | -101.9853
6.5601 | -101.9426
6.5638 | -101.9052
6.5675 | -101.8721
6.5712 | -101.846
6.5748 | -101.8373
6.5785 | -101.8305
6.5822 | -101.83
6.5859 | -101.8382
6.5895 | -101.8493
6.5932 | -101.8608
6.5969 | -101.8727
6.6006 | -101.8845
6.6042 | -101.8927
6.6079 | -101.888
6.6116 | -101.8816
6.6152 | -101.8815
6.6189 | -101.8828
6.6226 | -101.8849
6.6263 | -101.8872
6.6299 | -101.8967
6.6336 | -101.9056
6.6373 | -101.9029
6.641 | -101.8974
6.6446 | -101.8896
6.6483 | -101.8749
6.652 | -101.8504
6.6556 | -101.8254
6.6593 | -101.798
6.663 | -101.7652
6.6667 | -101.7349
6.6703 | -101.7026
6.674 | -101.6685
6.6777 | -101.6397
6.6814 | -101.6108
6.685 | -101.5867
6.6887 | -101.5651
6.6924 | -101.5502
6.6961 | -101.5438
6.6997 | -101.5459
6.7034 | -101.5557
6.7071 | -101.5755
6.7107 | -101.6018
6.7144 | -101.629
6.7181 | -101.6631
6.7218 | -101.6994
6.7254 | -101.7375
6.7291 | -101.7749
6.7328 | -101.8173
6.7365 | -101.8703
6.7401 | -101.9224
6.7438 | -101.9754
6.7475 | -102.0364
6.7511 | -102.1005
6.7548 | -102.1567
6.7585 | -102.2056
6.7622 | -102.2483
6.7658 | -102.2835
6.7695 | -102.3104
6.7732 | -102.3265
6.7769 | -102.3266
6.7805 | -102.316
6.7842 | -102.2932
6.7879 | -102.2702
6.7916 | -102.2395
6.7952 | -102.202
6.7989 | -102.168
6.8026 | -102.1378
6.8062 | -102.1106
6.8099 | -102.0866
6.8136 | -102.0713
6.8173 | -102.0606
6.8209 | -102.0584
6.8246 | -102.0672
6.8283 | -102.0796
6.832 | -102.0959
6.8356 | -102.1171
6.8393 | -102.1417
6.843 | -102.1711
6.8466 | -102.197
6.8503 | -102.2245
6.854 | -102.2533
6.8577 | -102.2864
6.8613 | -102.3148
6.865 | -102.3269
6.8687 | -102.3345
6.8724 | -102.3357
6.876 | -102.3312
6.8797 | -102.3142
6.8834 | -102.2955
6.8871 | -102.2689
6.8907 | -102.2486
6.8944 | -102.2291
6.8981 | -102.2143
6.9017 | -102.2021
6.9054 | -102.1979
6.9091 | -102.1954
6.9128 | -102.1953
6.9164 | -102.1982
6.9201 | -102.2015
6.9238 | -102.1984
6.9275 | -102.1932
6.9311 | -102.1895
6.9348 | -102.1874
6.9385 | -102.1773
6.9421 | -102.1605
6.9458 | -102.1404
6.9495 | -102.1176
6.9532 | -102.0975
6.9568 | -102.0797
6.9605 | -102.062
6.9642 | -102.0517
6.9679 | -102.0491
6.9715 | -102.0416
6.9752 | -102.0303
6.9789 | -102.0303
6.9826 | -102.038
6.9862 | -102.0498
6.9899 | -102.0728
6.9936 | -102.1063
6.9972 | -102.1443
7.0009 | -102.1936
7.0046 | -102.2445
7.0083 | -102.3018
7.0119 | -102.3613
7.0156 | -102.4291
7.0193 | -102.5009
7.023 | -102.5708
7.0266 | -102.6375
7.0303 | -102.7015
7.034 | -102.7543
7.0376 | -102.7991
7.0413 | -102.839
7.045 | -102.8734
7.0487 | -102.8972
7.0523 | -102.9038
7.056 | -102.8976
7.0597 | -102.8757
7.0634 | -102.8454
7.067 | -102.8042
7.0707 | -102.7562
7.0744 | -102.7101
7.0781 | -102.6681
7.0817 | -102.6289
7.0854 | -102.5845
7.0891 | -102.5526
7.0927 | -102.5272
7.0964 | -102.5032
7.1001 | -102.4841
7.1038 | -102.4765
7.1074 | -102.4712
7.1111 | -102.4701
7.1148 | -102.4648
7.1185 | -102.4622
7.1221 | -102.4627
7.1258 | -102.4608
7.1295 | -102.4633
7.1331 | -102.4738
7.1368 | -102.4889
7.1405 | -102.5076
7.1442 | -102.5236
7.1478 | -102.5341
7.1515 | -102.5435
7.1552 | -102.5522
7.1589 | -102.5598
7.1625 | -102.5595
7.1662 | -102.5447
7.1699 | -102.5224
7.1736 | -102.4998
7.1772 | -102.477
7.1809 | -102.4478
7.1846 | -102.4143
7.1882 | -102.3804
7.1919 | -102.3445
7.1956 | -102.3064
7.1993 | -102.2722
7.2029 | -102.2453
7.2066 | -102.2263
7.2103 | -102.2114
7.214 | -102.2124
7.2176 | -102.2285
7.2213 | -102.2541
7.225 | -102.287
7.2287 | -102.3256
7.2323 | -102.3737
7.236 | -102.4294
7.2397 | -102.4861
7.2433 | -102.5417
7.247 | -102.5991
7.2507 | -102.6523
7.2544 | -102.7022
7.258 | -102.7467
7.2617 | -102.7829
7.2654 | -102.8082
7.2691 | -102.8304
7.2727 | -102.8481
7.2764 | -102.853
7.2801 | -102.8452
7.2837 | -102.832
7.2874 | -102.8158
7.2911 | -102.7951
7.2948 | -102.7758
7.2984 | -102.7588
7.3021 | -102.7378
7.3058 | -102.719
7.3095 | -102.7023
7.3131 | -102.6911
7.3168 | -102.6839
7.3205 | -102.682
7.3242 | -102.6924
7.3278 | -102.7141
7.3315 | -102.7489
7.3352 | -102.7866
7.3388 | -102.8264
7.3425 | -102.8689
7.3462 | -102.9142
7.3499 | -102.9575
7.3535 | -102.9941
7.3572 | -103.03
7.3609 | -103.0548
7.3646 | -103.0654
7.3682 | -103.0604
7.3719 | -103.0471
7.3756 | -103.0188
7.3792 | -102.9758
7.3829 | -102.9208
7.3866 | -102.8596
7.3903 | -102.7971
7.3939 | -102.7341
7.3976 | -102.6738
7.4013 | -102.6181
7.405 | -102.5701
7.4086 | -102.533
7.4123 | -102.5018
7.416 | -102.5097
7.4197 | -102.5049
7.4233 | -102.5095
7.427 | -102.5218
7.4307 | -102.5374
7.4343 | -102.5588
7.438 | -102.5793
7.4417 | -102.5932
7.4454 | -102.6088
7.449 | -102.6303
7.4527 | -102.6497
7.4564 | -102.6638
7.4601 | -102.6732
7.4637 | -102.6818
7.4674 | -102.6826
7.4711 | -102.6851
7.4747 | -102.6935
7.4784 | -102.7044
7.4821 | -102.7108
7.4858 | -102.7198
7.4894 | -102.7311
7.4931 | -102.7427
7.4968 | -102.7501
7.5005 | -102.7528
7.5041 | -102.7571
7.5078 | -102.7584
7.5115 | -102.7554
7.5152 | -102.7535
7.5188 | -102.7505
7.5225 | -102.7437
7.5262 | -102.733
7.5298 | -102.7209
7.5335 | -102.7082
7.5372 | -102.6964
7.5409 | -102.6856
7.5445 | -102.6757
7.5482 | -102.6615
7.5519 | -102.6556
7.5556 | -102.6539
7.5592 | -102.6533
7.5629 | -102.6499
7.5666 | -102.6493
7.5702 | -102.6536
7.5739 | -102.6641
7.5776 | -102.6683
7.5813 | -102.6731
7.5849 | -102.6778
7.5886 | -102.6781
7.5923 | -102.6809
7.596 | -102.6853
7.5996 | -102.6875
7.6033 | -102.6845
7.607 | -102.6818
7.6107 | -102.6801
7.6143 | -102.6796
7.618 | -102.6778
7.6217 | -102.6809
7.6253 | -102.6876
7.629 | -102.6963
7.6327 | -102.7126
7.6364 | -102.7341
7.64 | -102.7556
7.6437 | -102.7799
7.6474 | -102.8006
7.6511 | -102.82
7.6547 | -102.8417
7.6584 | -102.8665
7.6621 | -102.8905
7.6657 | -102.9065
7.6694 | -102.9191
7.6731 | -102.9267
7.6768 | -102.9261
7.6804 | -102.9199
7.6841 | -102.9116
7.6878 | -102.8981
7.6915 | -102.8821
7.6951 | -102.8604
7.6988 | -102.8305
7.7025 | -102.8045
7.7062 | -102.7807
7.7098 | -102.7599
7.7135 | -102.74
7.7172 | -102.7226
7.7208 | -102.7068
7.7245 | -102.6927
7.7282 | -102.6823
7.7319 | -102.6783
7.7355 | -102.6737
7.7392 | -102.6731
7.7429 | -102.6736
7.7466 | -102.6767
7.7502 | -102.6834
7.7539 | -102.6945
7.7576 | -102.7084
7.7612 | -102.7325
7.7649 | -102.7594
7.7686 | -102.7858
7.7723 | -102.8109
7.7759 | -102.8364
7.7796 | -102.8554
7.7833 | -102.8683
7.787 | -102.8773
7.7906 | -102.8812
7.7943 | -102.8813
7.798 | -102.8734
7.8017 | -102.8607
7.8053 | -102.8491
7.809 | -102.842
7.8127 | -102.8323
7.8163 | -102.8185
7.82 | -102.8106
7.8237 | -102.8104
7.8274 | -102.8195
7.831 | -102.8366
7.8347 | -102.8568
7.8384 | -102.8833
7.8421 | -102.906
7.8457 | -102.9274
7.8494 | -102.9534
7.8531 | -102.9752
7.8567 | -102.9909
7.8604 | -102.9999
7.8641 | -103.0078
7.8678 | -103.0061
7.8714 | -102.9974
7.8751 | -102.9798
7.8788 | -102.9569
7.8825 | -102.9308
7.8861 | -102.9035
7.8898 | -102.8788
7.8935 | -102.8612
7.8972 | -102.8452
7.9008 | -102.8311
7.9045 | -102.8257
7.9082 | -102.8284
7.9118 | -102.8297
7.9155 | -102.8363
7.9192 | -102.8477
7.9229 | -102.8593
7.9265 | -102.8796
7.9302 | -102.9024
7.9339 | -102.9215
7.9376 | -102.9367
7.9412 | -102.9445
7.9449 | -102.9194
7.9486 | -102.8729
7.9523 | -102.8456
7.9559 | -102.8107
7.9596 | -102.7723
7.9633 | -102.7262
7.9669 | -102.6705
7.9706 | -102.6107
7.9743 | -102.552
7.978 | -102.4948
7.9816 | -102.4386
7.9853 | -102.3906
7.989 | -102.3523
7.9927 | -102.3269
7.9963 | -102.3166
8 | -102.3198
8.0037 | -102.3387
8.0073 | -102.3691
8.011 | -102.413
8.0147 | -102.4693
8.0184 | -102.5366
8.022 | -102.6073
8.0257 | -102.6774
8.0294 | -102.7503
8.0331 | -102.8163
8.0367 | -102.8777
8.0404 | -102.9287
8.0441 | -102.9699
8.0478 | -103.0001
8.0514 | -103.0249
8.0551 | -103.041
8.0588 | -103.052
8.0624 | -103.078
8.0661 | -103.0744
8.0698 | -103.0624
8.0735 | -103.0537
8.0771 | -103.0389
8.0808 | -103.0205
8.0845 | -103.0018
8.0882 | -102.9858
8.0918 | -102.9728
8.0955 | -102.964
8.0992 | -102.9606
8.1028 | -102.9575
8.1065 | -102.9574
8.1102 | -102.9615
8.1139 | -102.9699
8.1175 | -102.9808
8.1212 | -102.9957
8.1249 | -103.009
8.1286 | -103.0243
8.1322 | -103.0389
8.1359 | -103.0506
8.1396 | -103.0599
8.1433 | -103.0714
8.1469 | -103.0822
8.1506 | -103.0941
8.1543 | -103.1012
8.1579 | -103.1006
8.1616 | -103.1016
8.1653 | -103.0979
8.169 | -103.0959
8.1726 | -103.0909
8.1763 | -103.0906
8.18 | -103.0883
8.1837 | -103.0807
8.1873 | -103.0699
8.191 | -103.0596
8.1947 | -103.0424
8.1983 | -103.0245
8.202 | -103.0118
8.2057 | -103.0009
8.2094 | -102.989
8.213 | -102.9698
8.2167 | -102.9505
8.2204 | -102.9292
8.2241 | -102.9109
8.2277 | -102.8877
8.2314 | -102.8636
8.2351 | -102.8372
8.2388 | -102.8095
8.2424 | -102.782
8.2461 | -102.7569
8.2498 | -102.7341
8.2534 | -102.7169
8.2571 | -102.7031
8.2608 | -102.6876
8.2645 | -102.6734
8.2681 | -102.6609
8.2718 | -102.6507
8.2755 | -102.6482
8.2792 | -102.6526
8.2828 | -102.6597
8.2865 | -102.6709
8.2902 | -102.6834
8.2938 | -102.6966
8.2975 | -102.7124
8.3012 | -102.7372
8.3049 | -102.7692
8.3085 | -102.8086
8.3122 | -102.8449
8.3159 | -102.887
8.3196 | -102.9285
8.3232 | -102.9613
8.3269 | -102.9923
8.3306 | -103.0211
8.3343 | -103.048
8.3379 | -103.0666
8.3416 | -103.082
8.3453 | -103.0943
8.3489 | -103.1049
8.3526 | -103.1042
8.3563 | -103.1002
8.36 | -103.0957
8.3636 | -103.0898
8.3673 | -103.0805
8.371 | -103.0696
8.3747 | -103.0568
8.3783 | -103.0413
8.382 | -103.0319
8.3857 | -103.0247
8.3893 | -103.0198
8.393 | -103.0151
8.3967 | -103.0111
8.4004 | -103.0059
8.404 | -103.0009
8.4077 | -103.0018
8.4114 | -103.0028
8.4151 | -103.0059
8.4187 | -103.0099
8.4224 | -103.0108
8.4261 | -103.0126
8.4298 | -103.0168
8.4334 | -103.0277
8.4371 | -103.0363
8.4408 | -103.0469
8.4444 | -103.0612
8.4481 | -103.0755
8.4518 | -103.0884
8.4555 | -103.1037
8.4591 | -103.1229
8.4628 | -103.1333
8.4665 | -103.1387
8.4702 | -103.149
8.4738 | -103.1588
8.4775 | -103.1498
8.4812 | -103.151
8.4848 | -103.1454
8.4885 | -103.1324
8.4922 | -103.1199
8.4959 | -103.101
8.4995 | -103.0821
8.5032 | -103.0561
8.5069 | -103.0208
8.5106 | -102.9886
8.5142 | -102.9563
8.5179 | -102.9212
8.5216 | -102.8825
8.5253 | -102.8417
8.5289 | -102.8037
8.5326 | -102.7617
8.5363 | -102.7242
8.5399 | -102.6913
8.5436 | -102.6626
8.5473 | -102.6409
8.551 | -102.6251
8.5546 | -102.6149
8.5583 | -102.6104
8.562 | -102.6167
8.5657 | -102.6361
8.5693 | -102.6617
8.573 | -102.6952
8.5767 | -102.7396
8.5803 | -102.79
8.584 | -102.8422
8.5877 | -102.9068
8.5914 | -102.9729
8.595 | -103.028
8.5987 | -103.0824
8.6024 | -103.14
8.6061 | -103.1948
8.6097 | -103.2421
8.6134 | -103.2856
8.6171 | -103.3259
8.6208 | -103.3552
8.6244 | -103.3825
8.6281 | -103.4023
8.6318 | -103.4227
8.6354 | -103.4391
8.6391 | -103.4395
8.6428 | -103.4397
8.6465 | -103.4443
8.6501 | -103.448
8.6538 | -103.4527
8.6575 | -103.4575
8.6612 | -103.4665
8.6648 | -103.4716
8.6685 | -103.472
8.6722 | -103.4708
8.6758 | -103.4684
8.6795 | -103.4648
8.6832 | -103.4569
8.6869 | -103.4418
8.6905 | -103.4222
8.6942 | -103.3898
8.6979 | -103.3561
8.7016 | -103.3159
8.7052 | -103.2696
8.7089 | -103.2217
8.7126 | -103.1737
8.7163 | -103.1252
8.7199 | -103.0796
8.7236 | -103.0429
8.7273 | -103.0122
8.7309 | -102.9826
8.7346 | -102.9569
8.7383 | -102.9377
8.742 | -102.9225
8.7456 | -102.9154
8.7493 | -102.9156
8.753 | -102.9209
8.7567 | -102.9294
8.7603 | -102.9448
8.764 | -102.9636
8.7677 | -102.9918
8.7713 | -103.0185
8.775 | -103.0438
8.7787 | -103.0706
8.7824 | -103.1034
8.786 | -103.1337
8.7897 | -103.1599
8.7934 | -103.1835
8.7971 | -103.2014
8.8007 | -103.2189
8.8044 | -103.2322
8.8081 | -103.24
8.8118 | -103.248
8.8154 | -103.2567
8.8191 | -103.2629
8.8228 | -103.2726
8.8264 | -103.2809
8.8301 | -103.2942
8.8338 | -103.3074
8.8375 | -103.3241
8.8411 | -103.3444
8.8448 | -103.3631
8.8485 | -103.3859
8.8522 | -103.4148
8.8558 | -103.4453
8.8595 | -103.4799
8.8632 | -103.5131
8.8668 | -103.5516
8.8705 | -103.5955
8.8742 | -103.6307
8.8779 | -103.6627
8.8815 | -103.693
8.8852 | -103.7201
8.8889 | -103.7405
8.8926 | -103.7539
8.8962 | -103.7544
8.8999 | -103.7501
8.9036 | -103.7395
8.9073 | -103.7227
8.9109 | -103.6958
8.9146 | -103.668
8.9183 | -103.6388
8.9219 | -103.6096
8.9256 | -103.5861
8.9293 | -103.5583
8.933 | -103.5269
8.9366 | -103.4939
8.9403 | -103.4651
8.944 | -103.4384
8.9477 | -103.4076
8.9513 | -103.3794
8.955 | -103.3517
8.9587 | -103.32
8.9624 | -103.2857
8.966 | -103.2464
8.9697 | -103.2054
8.9734 | -103.1604
8.977 | -103.1139
8.9807 | -103.0638
8.9844 | -103.0167
8.9881 | -102.9758
8.9917 | -102.942
8.9954 | -102.9119
8.9991 | -102.8856
9.0028 | -102.8716
9.0064 | -102.8906
9.0101 | -102.8934
9.0138 | -102.9015
9.0174 | -102.9119
9.0211 | -102.9289
9.0248 | -102.9528
9.0285 | -102.9845
9.0321 | -103.0191
9.0358 | -103.0558
9.0395 | -103.0939
9.0432 | -103.1362
9.0468 | -103.1779
9.0505 | -103.2186
9.0542 | -103.2635
9.0579 | -103.3092
9.0615 | -103.3588
9.0652 | -103.4129
9.0689 | -103.4669
9.0725 | -103.5152
9.0762 | -103.5587
9.0799 | -103.6064
9.0836 | -103.6534
9.0872 | -103.6947
9.0909 | -103.7318
9.0946 | -103.7633
9.0983 | -103.7928
9.1019 | -103.8162
9.1056 | -103.8423
9.1093 | -103.8588
9.1129 | -103.8664
9.1166 | -103.8658
9.1203 | -103.8564
9.124 | -103.8392
9.1276 | -103.8161
9.1313 | -103.7895
9.135 | -103.7558
9.1387 | -103.7202
9.1423 | -103.6881
9.146 | -103.6588
9.1497 | -103.6307
9.1534 | -103.6159
9.157 | -103.6069
9.1607 | -103.6032
9.1644 | -103.6088
9.168 | -103.624
9.1717 | -103.6454
9.1754 | -103.6773
9.1791 | -103.7181
9.1827 | -103.7686
9.1864 | -103.8242
9.1901 | -103.8808
9.1938 | -103.9357
9.1974 | -103.9857
9.2011 | -104.0358
9.2048 | -104.0911
9.2084 | -104.1438
9.2121 | -104.1934
9.2158 | -104.2355
9.2195 | -104.2665
9.2231 | -104.2917
9.2268 | -104.3131
9.2305 | -104.3312
9.2342 | -104.3352
9.2378 | -104.3272
9.2415 | -104.3133
9.2452 | -104.2989
9.2489 | -104.2782
9.2525 | -104.25
9.2562 | -104.2163
9.2599 | -104.1829
9.2635 | -104.148
9.2672 | -104.1116
9.2709 | -104.0752
9.2746 | -104.0363
9.2782 | -103.999
9.2819 | -103.9624
9.2856 | -103.933
9.2893 | -103.9005
9.2929 | -103.8754
9.2966 | -103.8561
9.3003 | -103.8382
9.3039 | -103.8267
9.3076 | -103.8238
9.3113 | -103.8207
9.315 | -103.821
9.3186 | -103.8261
9.3223 | -103.8302
9.326 | -103.8299
9.3297 | -103.8378
9.3333 | -103.8494
9.337 | -103.8582
9.3407 | -103.8602
9.3444 | -103.8602
9.348 | -103.8581
9.3517 | -103.8526
9.3554 | -103.8457
9.359 | -103.838
9.3627 | -103.8265
9.3664 | -103.8115
9.3701 | -103.7993
9.3737 | -103.7845
9.3774 | -103.766
9.3811 | -103.7481
9.3848 | -103.7314
9.3884 | -103.7138
9.3921 | -103.7006
9.3958 | -103.6834
9.3994 | -103.6638
9.4031 | -103.6438
9.4068 | -103.6301
9.4105 | -103.6168
9.4141 | -103.6064
9.4178 | -103.5981
9.4215 | -103.5913
9.4252 | -103.5882
9.4288 | -103.5836
9.4325 | -103.5787
9.4362 | -103.5761
9.4399 | -103.5777
9.4435 | -103.5793
9.4472 | -103.5855
9.4509 | -103.5915
9.4545 | -103.5983
9.4582 | -103.6127
9.4619 | -103.6341
9.4656 | -103.6473
9.4692 | -103.66
9.4729 | -103.6754
9.4766 | -103.6924
9.4803 | -103.7057
9.4839 | -103.7186
9.4876 | -103.7303
9.4913 | -103.7419
9.4949 | -103.752
9.4986 | -103.7565
9.5023 | -103.7613
9.506 | -103.7646
9.5096 | -103.7678
9.5133 | -103.7687
9.517 | -103.7721
9.5207 | -103.7817
9.5243 | -103.796
9.528 | -103.8056
9.5317 | -103.8182
9.5354 | -103.834
9.539 | -103.8497
9.5427 | -103.8676
9.5464 | -103.8822
9.55 | -103.8982
9.5537 | -103.9146
9.5574 | -103.9299
9.5611 | -103.9393
9.5647 | -103.9497
9.5684 | -103.9619
9.5721 | -103.9723
9.5758 | -103.9774
9.5794 | -103.9769
9.5831 | -103.9744
9.5868 | -103.9703
9.5905 | -103.9656
9.5941 | -103.9598
9.5978 | -103.95
9.6015 | -103.9381
9.6051 | -103.9273
9.6088 | -103.9206
9.6125 | -103.9156
9.6162 | -103.9091
9.6198 | -103.905
9.6235 | -103.9029
9.6272 | -103.8973
9.6309 | -103.8919
9.6345 | -103.8828
9.6382 | -103.8735
9.6419 | -103.8654
9.6455 | -103.8569
9.6492 | -103.8444
9.6529 | -103.8298
9.6566 | -103.8129
9.6602 | -103.7973
9.6639 | -103.7796
9.6676 | -103.7673
9.6713 | -103.7544
9.6749 | -103.741
9.6786 | -103.7302
9.6823 | -103.7189
9.686 | -103.7051
9.6896 | -103.6939
9.6933 | -103.6874
9.697 | -103.6792
9.7006 | -103.6715
9.7043 | -103.6646
9.708 | -103.6598
9.7117 | -103.656
9.7153 | -103.6552
9.719 | -103.6598
9.7227 | -103.6676
9.7264 | -103.6741
9.73 | -103.6751
9.7337 | -103.6796
9.7374 | -103.6867
9.741 | -103.6967
9.7447 | -103.7018
9.7484 | -103.7097
9.7521 | -103.7121
9.7557 | -103.712
9.7594 | -103.7133
9.7631 | -103.7155
9.7668 | -103.715
9.7704 | -103.7147
9.7741 | -103.7124
9.7778 | -103.7159
9.7815 | -103.7186
9.7851 | -103.7146
9.7888 | -103.7078
9.7925 | -103.7077
9.7961 | -103.7067
9.7998 | -103.7028
9.8035 | -103.6958
9.8072 | -103.6921
9.8108 | -103.6885
9.8145 | -103.6844
9.8182 | -103.6787
9.8219 | -103.6724
9.8255 | -103.6686
9.8292 | -103.6698
9.8329 | -103.6713
9.8365 | -103.6748
9.8402 | -103.6824
9.8439 | -103.6927
9.8476 | -103.7085
9.8512 | -103.7193
9.8549 | -103.7296
9.8586 | -103.7416
9.8623 | -103.7496
9.8659 | -103.7625
9.8696 | -103.7758
9.8733 | -103.7909
9.877 | -103.8007
9.8806 | -103.8112
9.8843 | -103.82
9.888 | -103.8304
9.8916 | -103.8366
9.8953 | -103.843
9.899 | -103.8502
9.9027 | -103.8556
9.9063 | -103.8592
9.91 | -103.8613
9.9137 | -103.8641
9.9174 | -103.8621
9.921 | -103.8519
9.9247 | -103.8429
9.9284 | -103.8233
9.932 | -103.798
9.9357 | -103.7738
9.9394 | -103.7455
9.9431 | -103.7166
9.9467 | -103.687
9.9504 | -103.6531
9.9541 | -103.6129
9.9578 | -103.5752
9.9614 | -103.5363
9.9651 | -103.4907
9.9688 | -103.4424
9.9725 | -103.3918
9.9761 | -103.3447
9.9798 | -103.2984
9.9835 | -103.2555
9.9871 | -103.2195
9.9908 | -103.1851
9.9945 | -103.1563
9.9982 | -103.135
10.0018 | -103.1228
10.0055 | -103.1192
10.0092 | -103.1231
10.0129 | -103.1315
10.0165 | -103.1468
10.0202 | -103.1732
10.0239 | -103.2037
10.0275 | -103.2386
10.0312 | -103.2764
10.0349 | -103.3188
10.0386 | -103.3609
10.0422 | -103.4016
10.0459 | -103.444
10.0496 | -103.484
10.0533 | -103.52
10.0569 | -103.5521
10.0606 | -103.5828
10.0643 | -103.6133
10.068 | -103.6358
10.0716 | -103.6614
10.0753 | -103.698
10.079 | -103.7321
10.0826 | -103.7541
10.0863 | -103.7767
10.09 | -103.7983
10.0937 | -103.8208
10.0973 | -103.8467
10.101 | -103.8764
10.1047 | -103.9045
10.1084 | -103.9358
10.112 | -103.9677
10.1157 | -103.9978
10.1194 | -104.0283
10.123 | -104.0539
10.1267 | -104.078
10.1304 | -104.0992
10.1341 | -104.1233
10.1377 | -104.1478
10.1414 | -104.1678
10.1451 | -104.1807
10.1488 | -104.1924
10.1524 | -104.2039
10.1561 | -104.2071
10.1598 | -104.2187
10.1635 | -104.2256
10.1671 | -104.2223
10.1708 | -104.2148
10.1745 | -104.2059
10.1781 | -104.196
10.1818 | -104.1828
10.1855 | -104.1681
10.1892 | -104.1523
10.1928 | -104.1325
10.1965 | -104.1142
10.2002 | -104.0993
10.2039 | -104.0877
10.2075 | -104.0764
10.2112 | -104.0643
10.2149 | -104.0491
10.2185 | -104.041
10.2222 | -104.0387
10.2259 | -104.0319
10.2296 | -104.029
10.2332 | -104.0318
10.2369 | -104.0335
10.2406 | -104.0344
10.2443 | -104.0338
10.2479 | -104.0329
10.2516 | -104.0392
10.2553 | -104.044
10.259 | -104.0467
10.2626 | -104.0534
10.2663 | -104.0572
10.27 | -104.0546
10.2736 | -104.0472
10.2773 | -104.0405
10.281 | -104.0324
10.2847 | -104.0134
10.2883 | -103.9966
10.292 | -103.9828
10.2957 | -103.967
10.2994 | -103.9469
10.303 | -103.9271
10.3067 | -103.9096
10.3104 | -103.8897
10.314 | -103.8721
10.3177 | -103.8579
10.3214 | -103.8536
10.3251 | -103.8538
10.3287 | -103.8573
10.3324 | -103.8601
10.3361 | -103.8668
10.3398 | -103.877
10.3434 | -103.8893
10.3471 | -103.9001
10.3508 | -103.9087
10.3545 | -103.9168
10.3581 | -103.9204
10.3618 | -103.9251
10.3655 | -103.931
10.3691 | -103.9303
10.3728 | -103.9256
10.3765 | -103.9242
10.3802 | -103.9189
10.3838 | -103.9098
10.3875 | -103.8985
10.3912 | -103.888
10.3949 | -103.8766
10.3985 | -103.8646
10.4022 | -103.8474
10.4059 | -103.8365
10.4095 | -103.8265
10.4132 | -103.8186
10.4169 | -103.8179
10.4206 | -103.8219
10.4242 | -103.8322
10.4279 | -103.8455
10.4316 | -103.8634
10.4353 | -103.8852
10.4389 | -103.9078
10.4426 | -103.9351
10.4463 | -103.9695
10.45 | -104.005
10.4536 | -104.042
10.4573 | -104.0821
10.461 | -104.1235
10.4646 | -104.164
10.4683 | -104.2008
10.472 | -104.2333
10.4757 | -104.2649
10.4793 | -104.2935
10.483 | -104.3234
10.4867 | -104.3497
10.4904 | -104.3716
10.494 | -104.3935
10.4977 | -104.4136
10.5014 | -104.432
10.5051 | -104.4437
10.5087 | -104.4608
10.5124 | -104.4701
10.5161 | -104.4737
10.5197 | -104.4761
10.5234 | -104.4757
10.5271 | -104.4703
10.5308 | -104.4667
10.5344 | -104.4582
10.5381 | -104.4474
10.5418 | -104.4367
10.5455 | -104.4237
10.5491 | -104.409
10.5528 | -104.3929
10.5565 | -104.3721
10.5601 | -104.3547
10.5638 | -104.337
10.5675 | -104.316
10.5712 | -104.2961
10.5748 | -104.2768
10.5785 | -104.2594
10.5822 | -104.247
10.5859 | -104.2401
10.5895 | -104.2304
10.5932 | -104.2154
10.5969 | -104.2099
10.6006 | -104.1942
10.6042 | -104.1826
10.6079 | -104.1742
10.6116 | -104.1657
10.6152 | -104.1548
10.6189 | -104.1449
10.6226 | -104.1367
10.6263 | -104.127
10.6299 | -104.1146
10.6336 | -104.1037
10.6373 | -104.095
10.641 | -104.0875
10.6446 | -104.0798
10.6483 | -104.0692
10.652 | -104.0608
10.6556 | -104.0554
10.6593 | -104.052
10.663 | -104.0487
10.6667 | -104.05
10.6703 | -104.0526
10.674 | -104.055
10.6777 | -104.0613
10.6814 | -104.0711
10.685 | -104.0796
10.6887 | -104.092
10.6924 | -104.1082
10.6961 | -104.121
10.6997 | -104.1352
10.7034 | -104.1486
10.7071 | -104.161
10.7107 | -104.1757
10.7144 | -104.1859
10.7181 | -104.193
10.7218 | -104.2016
10.7254 | -104.2063
10.7291 | -104.2099
10.7328 | -104.2078
10.7365 | -104.2009
10.7401 | -104.1959
10.7438 | -104.1946
10.7475 | -104.1962
10.7511 | -104.188
10.7548 | -104.1772
10.7585 | -104.172
10.7622 | -104.1655
10.7658 | -104.156
10.7695 | -104.1479
10.7732 | -104.1423
10.7769 | -104.1369
10.7805 | -104.1327
10.7842 | -104.1285
10.7879 | -104.1261
10.7916 | -104.1212
10.7952 | -104.1164
10.7989 | -104.1125
10.8026 | -104.11
10.8062 | -104.1088
10.8099 | -104.1092
10.8136 | -104.1089
10.8173 | -104.1128
10.8209 | -104.1167
10.8246 | -104.1219
10.8283 | -104.1291
10.832 | -104.1405
10.8356 | -104.1498
10.8393 | -104.1565
10.843 | -104.1692
10.8466 | -104.1878
10.8503 | -104.2069
10.854 | -104.2274
10.8577 | -104.248
10.8613 | -104.2719
10.865 | -104.2976
10.8687 | -104.3202
10.8724 | -104.3393
10.876 | -104.3566
10.8797 | -104.373
10.8834 | -104.3869
10.8871 | -104.3969
10.8907 | -104.4083
10.8944 | -104.4184
10.8981 | -104.422
10.9017 | -104.4195
10.9054 | -104.4136
10.9091 | -104.4115
10.9128 | -104.4097
10.9164 | -104.4018
10.9201 | -104.3886
10.9238 | -104.3683
10.9275 | -104.3496
10.9311 | -104.3285
10.9348 | -104.3114
10.9385 | -104.2978
10.9421 | -104.2819
10.9458 | -104.2666
10.9495 | -104.2539
10.9532 | -104.2425
10.9568 | -104.2317
10.9605 | -104.2175
10.9642 | -104.2034
10.9679 | -104.1866
10.9715 | -104.1748
10.9752 | -104.1647
10.9789 | -104.1534
10.9826 | -104.1415
10.9862 | -104.1331
10.9899 | -104.1237
10.9936 | -104.1173
10.9972 | -104.1131
11.0009 | -104.1056
11.0046 | -104.0938
11.0083 | -104.0848
11.0119 | -104.0806
11.0156 | -104.0803
11.0193 | -104.0782
11.023 | -104.0788
11.0266 | -104.0819
11.0303 | -104.0899
11.034 | -104.0993
11.0376 | -104.1086
11.0413 | -104.1218
11.045 | -104.1377
11.0487 | -104.1538
11.0523 | -104.1746
11.056 | -104.1985
11.0597 | -104.2281
11.0634 | -104.2618
11.067 | -104.2925
11.0707 | -104.3263
11.0744 | -104.3644
11.0781 | -104.4047
11.0817 | -104.4479
11.0854 | -104.4861
11.0891 | -104.5217
11.0927 | -104.5545
11.0964 | -104.582
11.1001 | -104.6105
11.1038 | -104.6364
11.1074 | -104.6571
11.1111 | -104.6708
11.1148 | -104.6821
11.1185 | -104.6938
11.1221 | -104.7015
11.1258 | -104.6904
11.1295 | -104.6942
11.1332 | -104.6887
11.1368 | -104.6797
11.1405 | -104.6754
11.1442 | -104.6719
11.1478 | -104.6672
11.1515 | -104.6647
11.1552 | -104.6595
11.1589 | -104.6551
11.1625 | -104.6543
11.1662 | -104.6539
11.1699 | -104.6537
11.1736 | -104.6529
11.1772 | -104.6547
11.1809 | -104.6566
11.1846 | -104.6578
11.1882 | -104.6602
11.1919 | -104.6632
11.1956 | -104.6634
11.1993 | -104.6631
11.2029 | -104.6576
11.2066 | -104.6511
11.2103 | -104.6461
11.214 | -104.6417
11.2176 | -104.6392
11.2213 | -104.6324
11.225 | -104.6247
11.2287 | -104.6213
11.2323 | -104.611
11.236 | -104.5988
11.2397 | -104.5921
11.2433 | -104.585
11.247 | -104.5729
11.2507 | -104.5656
11.2544 | -104.5545
11.258 | -104.5424
11.2617 | -104.5274
11.2654 | -104.5152
11.2691 | -104.5032
11.2727 | -104.4901
11.2764 | -104.4765
11.2801 | -104.4642
11.2837 | -104.4534
11.2874 | -104.4415
11.2911 | -104.4299
11.2948 | -104.4187
11.2984 | -104.4142
11.3021 | -104.4106
11.3058 | -104.4065
11.3095 | -104.403
11.3131 | -104.4027
11.3168 | -104.4072
11.3205 | -104.4135
11.3242 | -104.4227
11.3278 | -104.4332
11.3315 | -104.4472
11.3352 | -104.4596
11.3388 | -104.4738
11.3425 | -104.4894
11.3462 | -104.5047
11.3499 | -104.5214
11.3535 | -104.5382
11.3572 | -104.5573
11.3609 | -104.5784
11.3646 | -104.5975
11.3682 | -104.6212
11.3719 | -104.644
11.3756 | -104.667
11.3792 | -104.6867
11.3829 | -104.7059
11.3866 | -104.7294
11.3903 | -104.7537
11.3939 | -104.7767
11.3976 | -104.8012
11.4013 | -104.8222
11.405 | -104.8367
11.4086 | -104.8534
11.4123 | -104.8681
11.416 | -104.8759
11.4197 | -104.8799
11.4233 | -104.8786
11.427 | -104.8742
11.4307 | -104.8676
11.4343 | -104.8588
11.438 | -104.8469
11.4417 | -104.8307
11.4454 | -104.8109
11.449 | -104.7886
11.4527 | -104.7685
11.4564 | -104.7424
11.4601 | -104.716
11.4637 | -104.6843
11.4674 | -104.6556
11.4711 | -104.6277
11.4747 | -104.599
11.4784 | -104.5733
11.4821 | -104.5524
11.4858 | -104.5357
11.4894 | -104.5186
11.4931 | -104.5022
11.4968 | -104.4884
11.5005 | -104.4792
11.5041 | -104.474
11.5078 | -104.4687
11.5115 | -104.4663
11.5152 | -104.4625
11.5188 | -104.4578
11.5225 | -104.4549
11.5262 | -104.4532
11.5298 | -104.4542
11.5335 | -104.4538
11.5372 | -104.4529
11.5409 | -104.4521
11.5445 | -104.4511
11.5482 | -104.4506
11.5519 | -104.449
11.5556 | -104.4475
11.5592 | -104.443
11.5629 | -104.4404
11.5666 | -104.4415
11.5702 | -104.4441
11.5739 | -104.4489
11.5776 | -104.4592
11.5813 | -104.4694
11.5849 | -104.4797
11.5886 | -104.4937
11.5923 | -104.5104
11.596 | -104.524
11.5996 | -104.5408
11.6033 | -104.5594
11.607 | -104.5822
11.6107 | -104.6054
11.6143 | -104.6249
11.618 | -104.6412
11.6217 | -104.6614
11.6253 | -104.6802
11.629 | -104.6972
11.6327 | -104.7068
11.6364 | -104.7127
11.64 | -104.7206
11.6437 | -104.7215
11.6474 | -104.7155
11.6511 | -104.7084
11.6547 | -104.6908
11.6584 | -104.682
11.6621 | -104.6676
11.6657 | -104.652
11.6694 | -104.6332
11.6731 | -104.611
11.6768 | -104.5868
11.6804 | -104.5615
11.6841 | -104.5367
11.6878 | -104.5133
11.6915 | -104.4877
11.6951 | -104.4626
11.6988 | -104.4386
11.7025 | -104.4187
11.7062 | -104.4025
11.7098 | -104.387
11.7135 | -104.3754
11.7172 | -104.3651
11.7208 | -104.3555
11.7245 | -104.3471
11.7282 | -104.3445
11.7319 | -104.345
11.7355 | -104.3428
11.7392 | -104.3402
11.7429 | -104.3379
11.7466 | -104.3378
11.7502 | -104.3393
11.7539 | -104.346
11.7576 | -104.3521
11.7612 | -104.3577
11.7649 | -104.3611
11.7686 | -104.3632
11.7723 | -104.3684
11.7759 | -104.3731
11.7796 | -104.3778
11.7833 | -104.3863
11.787 | -104.3949
11.7906 | -104.3984
11.7943 | -104.3995
11.798 | -104.4004
11.8017 | -104.4024
11.8053 | -104.4033
11.809 | -104.4021
11.8127 | -104.3951
11.8163 | -104.389
11.82 | -104.3854
11.8237 | -104.3811
11.8274 | -104.3779
11.831 | -104.3744
11.8347 | -104.3768
11.8384 | -104.3785
11.8421 | -104.3826
11.8457 | -104.3879
11.8494 | -104.3956
11.8531 | -104.4074
11.8567 | -104.4165
11.8604 | -104.4247
11.8641 | -104.4336
11.8678 | -104.4416
11.8714 | -104.4462
11.8751 | -104.4505
11.8788 | -104.4561
11.8825 | -104.4655
11.8861 | -104.4735
11.8898 | -104.4819
11.8935 | -104.4925
11.8972 | -104.5054
11.9008 | -104.5185
11.9045 | -104.5287
11.9082 | -104.5415
11.9118 | -104.552
11.9155 | -104.5537
11.9192 | -104.5604
11.9229 | -104.5713
11.9265 | -104.5803
11.9302 | -104.5871
11.9339 | -104.5929
11.9376 | -104.6001
11.9412 | -104.6018
11.9449 | -104.5972
11.9486 | -104.5853
11.9522 | -104.5728
11.9559 | -104.5558
11.9596 | -104.5337
11.9633 | -104.5103
11.9669 | -104.4846
11.9706 | -104.4569
11.9743 | -104.4243
11.978 | -104.3893
11.9816 | -104.3505
11.9853 | -104.3167
11.989 | -104.282
11.9927 | -104.2456
11.9963 | -104.2117
12 | -104.1821
12.0037 | -104.1569
12.0073 | -104.1338
12.011 | -104.1155
12.0147 | -104.1009
12.0184 | -104.0912
12.022 | -104.0855
12.0257 | -104.0879
12.0294 | -104.0987
12.0331 | -104.1134
12.0367 | -104.1295
12.0404 | -104.15
12.0441 | -104.1761
12.0478 | -104.2028
12.0514 | -104.2274
12.0551 | -104.2494
12.0588 | -104.2749
12.0624 | -104.3004
12.0661 | -104.3235
12.0698 | -104.3434
12.0735 | -104.359
12.0771 | -104.3709
12.0808 | -104.381
12.0845 | -104.3882
12.0882 | -104.3983
12.0918 | -104.3971
12.0955 | -104.3889
12.0992 | -104.3781
12.1028 | -104.3727
12.1065 | -104.3653
12.1102 | -104.3585
12.1139 | -104.3533
12.1175 | -104.3514
12.1212 | -104.3567
12.1249 | -104.3614
12.1286 | -104.368
12.1322 | -104.3789
12.1359 | -104.3952
12.1396 | -104.4143
12.1433 | -104.4364
12.1469 | -104.4631
12.1506 | -104.4937
12.1543 | -104.5223
12.1579 | -104.5517
12.1616 | -104.5836
12.1653 | -104.6174
12.169 | -104.6503
12.1726 | -104.6861
12.1763 | -104.7208
12.18 | -104.7526
12.1837 | -104.7801
12.1873 | -104.7925
12.191 | -104.8142
12.1947 | -104.8343
12.1983 | -104.8542
12.202 | -104.8665
12.2057 | -104.8773
12.2094 | -104.8859
12.213 | -104.8911
12.2167 | -104.8976
12.2204 | -104.901
12.2241 | -104.9023
12.2277 | -104.9006
12.2314 | -104.8979
12.2351 | -104.8976
12.2388 | -104.8975
12.2424 | -104.8947
12.2461 | -104.896
12.2498 | -104.8979
12.2534 | -104.9037
12.2571 | -104.906
12.2608 | -104.9074
12.2645 | -104.9107
12.2681 | -104.9146
12.2718 | -104.9187
12.2755 | -104.9247
12.2792 | -104.9268
12.2828 | -104.9289
12.2865 | -104.9299
12.2902 | -104.9313
12.2938 | -104.9283
12.2975 | -104.9245
12.3012 | -104.9193
12.3049 | -104.915
12.3085 | -104.909
12.3122 | -104.8998
12.3159 | -104.8886
12.3196 | -104.878
12.3232 | -104.8668
12.3269 | -104.8525
12.3306 | -104.8413
12.3343 | -104.8286
12.3379 | -104.8108
12.3416 | -104.7914
12.3453 | -104.7734
12.3489 | -104.7531
12.3526 | -104.7358
12.3563 | -104.7176
12.36 | -104.6995
12.3636 | -104.6801
12.3673 | -104.6622
12.371 | -104.6439
12.3747 | -104.6217
12.3783 | -104.6013
12.382 | -104.583
12.3857 | -104.5658
12.3893 | -104.549
12.393 | -104.5339
12.3967 | -104.5121
12.4004 | -104.4936
12.404 | -104.4733
12.4077 | -104.4571
12.4114 | -104.4413
12.4151 | -104.4208
12.4187 | -104.4021
12.4224 | -104.3872
12.4261 | -104.3705
12.4298 | -104.3583
12.4334 | -104.3452
12.4371 | -104.3311
12.4408 | -104.3209
12.4444 | -104.3083
12.4481 | -104.2957
12.4518 | -104.287
12.4555 | -104.2764
12.4591 | -104.2673
12.4628 | -104.2597
12.4665 | -104.2542
12.4702 | -104.2506
12.4738 | -104.2531
12.4775 | -104.2575
12.4812 | -104.2659
12.4848 | -104.2775
12.4885 | -104.2881
12.4922 | -104.3033
12.4959 | -104.3267
12.4995 | -104.3539
12.5032 | -104.3837
12.5069 | -104.414
12.5106 | -104.4486
12.5142 | -104.4816
12.5179 | -104.5174
12.5216 | -104.5517
12.5253 | -104.5792
12.5289 | -104.6056
12.5326 | -104.633
12.5363 | -104.6585
12.5399 | -104.6841
12.5436 | -104.7077
12.5473 | -104.7242
12.551 | -104.7367
12.5546 | -104.7449
12.5583 | -104.7472
12.562 | -104.7441
12.5657 | -104.739
12.5693 | -104.7344
12.573 | -104.7283
12.5767 | -104.7216
12.5803 | -104.7115
12.584 | -104.7043
12.5877 | -104.6958
12.5914 | -104.685
12.595 | -104.6739
12.5987 | -104.6653
12.6024 | -104.658
12.6061 | -104.654
12.6097 | -104.6512
12.6134 | -104.6529
12.6171 | -104.6581
12.6208 | -104.6658
12.6244 | -104.6769
12.6281 | -104.6888
12.6318 | -104.7004
12.6354 | -104.7179
12.6391 | -104.7387
12.6428 | -104.7604
12.6465 | -104.7795
12.6501 | -104.8007
12.6538 | -104.8208
12.6575 | -104.8377
12.6612 | -104.8531
12.6648 | -104.8691
12.6685 | -104.8885
12.6722 | -104.9038
12.6758 | -104.9118
12.6795 | -104.9177
12.6832 | -104.9233
12.6869 | -104.9319
12.6905 | -104.9324
12.6942 | -104.9321
12.6979 | -104.9297
12.7016 | -104.9268
12.7052 | -104.924
12.7089 | -104.9154
12.7126 | -104.9041
12.7163 | -104.893
12.7199 | -104.8807
12.7236 | -104.8694
12.7273 | -104.8558
12.7309 | -104.8458
12.7346 | -104.8379
12.7383 | -104.8297
12.742 | -104.8189
12.7456 | -104.8098
12.7493 | -104.7997
12.753 | -104.7926
12.7567 | -104.7871
12.7603 | -104.782
12.764 | -104.7746
12.7677 | -104.7665
12.7713 | -104.7592
12.775 | -104.7536
12.7787 | -104.7455
12.7824 | -104.7352
12.786 | -104.7267
12.7897 | -104.7163
12.7934 | -104.7048
12.7971 | -104.6949
12.8007 | -104.6876
12.8044 | -104.6792
12.8081 | -104.6705
12.8118 | -104.6646
12.8154 | -104.6579
12.8191 | -104.6489
12.8228 | -104.6367
12.8264 | -104.6327
12.8301 | -104.6305
12.8338 | -104.6256
12.8375 | -104.619
12.8411 | -104.6163
12.8448 | -104.6169
12.8485 | -104.6159
12.8522 | -104.6144
12.8558 | -104.6126
12.8595 | -104.6128
12.8632 | -104.6128
12.8668 | -104.6151
12.8705 | -104.6175
12.8742 | -104.6217
12.8779 | -104.6253
12.8815 | -104.6319
12.8852 | -104.6405
12.8889 | -104.6489
12.8926 | -104.6577
12.8962 | -104.6657
12.8999 | -104.6742
12.9036 | -104.6803
12.9073 | -104.6729
12.9109 | -104.6652
12.9146 | -104.663
12.9183 | -104.6594
12.9219 | -104.6511
12.9256 | -104.6449
12.9293 | -104.6366
12.933 | -104.6219
12.9366 | -104.6051
12.9403 | -104.5828
12.944 | -104.555
12.9477 | -104.5275
12.9513 | -104.5002
12.955 | -104.4707
12.9587 | -104.439
12.9624 | -104.4055
12.966 | -104.3723
12.9697 | -104.3402
12.9734 | -104.3042
12.977 | -104.2689
12.9807 | -104.2406
12.9844 | -104.2125
12.9881 | -104.1836
12.9917 | -104.1593
12.9954 | -104.1412
12.9991 | -104.1268
13.0028 | -104.1186
13.0064 | -104.113
13.0101 | -104.1165
13.0138 | -104.1207
13.0174 | -104.1266
13.0211 | -104.1378
13.0248 | -104.1496
13.0285 | -104.1624
13.0321 | -104.1764
13.0358 | -104.1924
13.0395 | -104.2074
13.0432 | -104.2244
13.0468 | -104.2409
13.0505 | -104.2562
13.0542 | -104.2698
13.0579 | -104.2824
13.0615 | -104.2952
13.0652 | -104.3056
13.0689 | -104.3109
13.0725 | -104.3125
13.0762 | -104.3121
13.0799 | -104.3139
13.0836 | -104.3148
13.0872 | -104.3146
13.0909 | -104.3168
13.0946 | -104.3275
13.0983 | -104.3297
13.1019 | -104.3244
13.1056 | -104.3188
13.1093 | -104.3145
13.1129 | -104.3083
13.1166 | -104.3018
13.1203 | -104.2989
13.124 | -104.293
13.1276 | -104.2868
13.1313 | -104.2826
13.135 | -104.2763
13.1387 | -104.2703
13.1423 | -104.264
13.146 | -104.257
13.1497 | -104.2489
13.1534 | -104.2407
13.157 | -104.231
13.1607 | -104.2202
13.1644 | -104.2108
13.168 | -104.1963
13.1717 | -104.1833
13.1754 | -104.1713
13.1791 | -104.1543
13.1827 | -104.1357
13.1864 | -104.1188
13.1901 | -104.1028
13.1938 | -104.0889
13.1974 | -104.0681
13.2011 | -104.0484
13.2048 | -104.036
13.2084 | -104.0215
13.2121 | -104.0104
13.2158 | -103.9992
13.2195 | -103.987
13.2231 | -103.9771
13.2268 | -103.9724
13.2305 | -103.9684
13.2342 | -103.9665
13.2378 | -103.9686
13.2415 | -103.9731
13.2452 | -103.9919
13.2489 | -104.002
13.2525 | -104.0158
13.2562 | -104.0335
13.2599 | -104.0528
13.2635 | -104.0764
13.2672 | -104.1046
13.2709 | -104.1329
13.2746 | -104.1627
13.2782 | -104.1927
13.2819 | -104.2258
13.2856 | -104.2607
13.2893 | -104.2944
13.2929 | -104.3273
13.2966 | -104.362
13.3003 | -104.3936
13.3039 | -104.4221
13.3076 | -104.449
13.3113 | -104.4733
13.315 | -104.496
13.3186 | -104.5164
13.3223 | -104.5295
13.326 | -104.538
13.3297 | -104.542
13.3333 | -104.5414
13.337 | -104.5382
13.3407 | -104.5324
13.3444 | -104.5235
13.348 | -104.5097
13.3517 | -104.497
13.3554 | -104.4787
13.359 | -104.4578
13.3627 | -104.4356
13.3664 | -104.4161
13.3701 | -104.3983
13.3737 | -104.3798
13.3774 | -104.3626
13.3811 | -104.345
13.3848 | -104.3296
13.3884 | -104.3203
13.3921 | -104.3129
13.3958 | -104.3058
13.3994 | -104.3017
13.4031 | -104.3
13.4068 | -104.3034
13.4105 | -104.3099
13.4141 | -104.3185
13.4178 | -104.3306
13.4215 | -104.3451
13.4252 | -104.3592
13.4288 | -104.3758
13.4325 | -104.3952
13.4362 | -104.4194
13.4399 | -104.444
13.4435 | -104.4652
13.4472 | -104.488
13.4509 | -104.513
13.4545 | -104.5389
13.4582 | -104.5572
13.4619 | -104.5738
13.4656 | -104.5898
13.4692 | -104.6018
13.4729 | -104.6125
13.4766 | -104.622
13.4803 | -104.6277
13.4839 | -104.6314
13.4876 | -104.6343
13.4913 | -104.632
13.4949 | -104.6286
13.4986 | -104.6239
13.5023 | -104.6139
13.506 | -104.6027
13.5096 | -104.5937
13.5133 | -104.5822
13.517 | -104.5727
13.5207 | -104.5628
13.5243 | -104.5508
13.528 | -104.5415
13.5317 | -104.5328
13.5354 | -104.5263
13.539 | -104.5225
13.5427 | -104.5226
13.5464 | -104.524
13.55 | -104.5268
13.5537 | -104.5328
13.5574 | -104.543
13.5611 | -104.5498
13.5647 | -104.56
13.5684 | -104.5756
13.5721 | -104.5923
13.5758 | -104.6107
13.5794 | -104.6303
13.5831 | -104.6518
13.5868 | -104.6742
13.5905 | -104.6964
13.5941 | -104.722
13.5978 | -104.7457
13.6015 | -104.7639
13.6051 | -104.781
13.6088 | -104.7975
13.6125 | -104.8092
13.6162 | -104.8203
13.6198 | -104.8351
13.6235 | -104.8453
13.6272 | -104.8527
13.6309 | -104.8565
13.6345 | -104.8597
13.6382 | -104.8625
13.6419 | -104.8634
13.6455 | -104.8627
13.6492 | -104.8611
13.6529 | -104.863
13.6566 | -104.8678
13.6602 | -104.8683
13.6639 | -104.8703
13.6676 | -104.8748
13.6713 | -104.876
13.6749 | -104.8769
13.6786 | -104.8825
13.6823 | -104.8899
13.686 | -104.8977
13.6896 | -104.9059
13.6933 | -104.9141
13.697 | -104.9256
13.7006 | -104.9338
13.7043 | -104.9389
13.708 | -104.9442
13.7117 | -104.952
13.7153 | -104.9592
13.719 | -104.9657
13.7227 | -104.9664
13.7264 | -104.97
13.73 | -104.9757
13.7337 | -104.9774
13.7374 | -104.9778
13.741 | -104.9789
13.7447 | -104.98
13.7484 | -104.9778
13.7521 | -104.9724
13.7557 | -104.9728
13.7594 | -104.9689
13.7631 | -104.958
13.7668 | -104.9452
13.7704 | -104.9351
13.7741 | -104.9231
13.7778 | -104.9095
13.7815 | -104.8925
13.7851 | -104.8717
13.7888 | -104.8537
13.7925 | -104.8367
13.7961 | -104.8178
13.7998 | -104.8018
13.8035 | -104.786
13.8072 | -104.7704
13.8108 | -104.7532
13.8145 | -104.7335
13.8182 | -104.7112
13.8219 | -104.6897
13.8255 | -104.6697
13.8292 | -104.6516
13.8329 | -104.6315
13.8365 | -104.6124
13.8402 | -104.5948
13.8439 | -104.5779
13.8476 | -104.5586
13.8512 | -104.5404
13.8549 | -104.5193
13.8586 | -104.5021
13.8623 | -104.4858
13.8659 | -104.4693
13.8696 | -104.4538
13.8733 | -104.4416
13.877 | -104.4279
13.8806 | -104.4141
13.8843 | -104.4052
13.888 | -104.3975
13.8916 | -104.391
13.8953 | -104.3857
13.899 | -104.3756
13.9027 | -104.3679
13.9063 | -104.3676
13.91 | -104.3671
13.9137 | -104.3689
13.9174 | -104.3723
13.921 | -104.3793
13.9247 | -104.3836
13.9284 | -104.3863
13.932 | -104.3899
13.9357 | -104.3974
13.9394 | -104.406
13.9431 | -104.4148
13.9467 | -104.4246
13.9504 | -104.4338
13.9541 | -104.4404
13.9578 | -104.4497
13.9614 | -104.463
13.9651 | -104.4778
13.9688 | -104.4916
13.9725 | -104.5031
13.9761 | -104.5178
13.9798 | -104.5367
13.9835 | -104.5544
13.9871 | -104.5716
13.9908 | -104.5837
13.9945 | -104.5955
13.9982 | -104.612
14.0018 | -104.6279
14.0055 | -104.6433
14.0092 | -104.6589
14.0129 | -104.6772
14.0165 | -104.6972
14.0202 | -104.7198
14.0239 | -104.7415
14.0275 | -104.7601
14.0312 | -104.7778
14.0349 | -104.7951
14.0386 | -104.8118
14.0422 | -104.8268
14.0459 | -104.84
14.0496 | -104.8579
14.0533 | -104.8762
14.0569 | -104.8943
14.0606 | -104.9131
14.0643 | -104.9308
14.068 | -104.9483
14.0716 | -104.9665
14.0753 | -104.9835
14.079 | -105.0018
14.0826 | -105.0178
14.0863 | -105.0331
14.09 | -105.0505
14.0937 | -105.0673
14.0973 | -105.0857
14.101 | -105.1123
14.1047 | -105.1367
14.1084 | -105.1501
14.112 | -105.1627
14.1157 | -105.1784
14.1194 | -105.1916
14.123 | -105.2063
14.1267 | -105.2203
14.1304 | -105.2325
14.1341 | -105.2485
14.1377 | -105.2573
14.1414 | -105.2648
14.1451 | -105.2759
14.1488 | -105.2841
14.1524 | -105.2915
14.1561 | -105.297
14.1598 | -105.2999
14.1635 | -105.3003
14.1671 | -105.2976
14.1708 | -105.2953
14.1745 | -105.2961
14.1781 | -105.2954
14.1818 | -105.2926
14.1855 | -105.2921
14.1892 | -105.2925
14.1928 | -105.2938
14.1965 | -105.2887
14.2002 | -105.2837
14.2039 | -105.279
14.2075 | -105.2731
14.2112 | -105.2679
14.2149 | -105.266
14.2185 | -105.2657
14.2222 | -105.2641
14.2259 | -105.264
14.2296 | -105.2656
14.2332 | -105.2688
14.2369 | -105.2731
14.2406 | -105.2743
14.2443 | -105.2756
14.2479 | -105.2772
14.2516 | -105.2745
14.2553 | -105.2712
14.259 | -105.269
14.2626 | -105.2628
14.2663 | -105.2571
14.27 | -105.2486
14.2736 | -105.2377
14.2773 | -105.2272
14.281 | -105.2156
14.2847 | -105.2032
14.2883 | -105.1893
14.292 | -105.174
14.2957 | -105.1558
14.2994 | -105.1369
14.303 | -105.1177
14.3067 | -105.0974
14.3104 | -105.0725
14.314 | -105.0513
14.3177 | -105.0271
14.3214 | -105.0029
14.3251 | -104.9829
14.3287 | -104.963
14.3324 | -104.9469
14.3361 | -104.9306
14.3398 | -104.9111
14.3434 | -104.8964
14.3471 | -104.8838
14.3508 | -104.8697
14.3545 | -104.8599
14.3581 | -104.8505
14.3618 | -104.846
14.3655 | -104.8434
14.3691 | -104.8408
14.3728 | -104.8405
14.3765 | -104.8385
14.3802 | -104.8334
14.3838 | -104.8324
14.3875 | -104.8366
14.3912 | -104.8428
14.3949 | -104.8485
14.3985 | -104.8534
14.4022 | -104.8602
14.4059 | -104.8636
14.4095 | -104.8673
14.4132 | -104.8729
14.4169 | -104.8788
14.4206 | -104.8838
14.4242 | -104.8874
14.4279 | -104.8915
14.4316 | -104.8971
14.4353 | -104.9027
14.4389 | -104.9077
14.4426 | -104.9123
14.4463 | -104.9207
14.45 | -104.9275
14.4536 | -104.935
14.4573 | -104.9421
14.461 | -104.9476
14.4646 | -104.9506
14.4683 | -104.9552
14.472 | -104.9624
14.4757 | -104.9707
14.4793 | -104.98
14.483 | -104.9872
14.4867 | -104.9934
14.4904 | -105.0017
14.494 | -105.0098
14.4977 | -105.017
14.5014 | -105.0247
14.5051 | -105.0319
14.5087 | -105.0418
14.5124 | -105.0491
14.5161 | -105.0588
14.5197 | -105.0747
14.5234 | -105.0923
14.5271 | -105.1084
14.5308 | -105.1215
14.5344 | -105.1341
14.5381 | -105.1455
14.5418 | -105.1569
14.5455 | -105.1704
14.5491 | -105.1838
14.5528 | -105.1971
14.5565 | -105.2084
14.5601 | -105.2209
14.5638 | -105.2306
14.5675 | -105.2404
14.5712 | -105.247
14.5748 | -105.2538
14.5785 | -105.2609
14.5822 | -105.2658
14.5859 | -105.2694
14.5895 | -105.2752
14.5932 | -105.2748
14.5969 | -105.2706
14.6006 | -105.2667
14.6042 | -105.2616
14.6079 | -105.2579
14.6116 | -105.2537
14.6152 | -105.2482
14.6189 | -105.2386
14.6226 | -105.2261
14.6263 | -105.2094
14.6299 | -105.194
14.6336 | -105.18
14.6373 | -105.1655
14.641 | -105.1522
14.6446 | -105.1414
14.6483 | -105.132
14.652 | -105.1232
14.6556 | -105.1116
14.6593 | -105.1016
14.663 | -105.0938
14.6667 | -105.0876
14.6703 | -105.0816
14.674 | -105.0787
14.6777 | -105.0767
14.6814 | -105.0734
14.685 | -105.0702
14.6887 | -105.069
14.6924 | -105.0719
14.6961 | -105.0758
14.6997 | -105.0793
14.7034 | -105.0849
14.7071 | -105.0923
14.7107 | -105.1017
14.7144 | -105.1109
14.7181 | -105.1159
14.7218 | -105.1267
14.7254 | -105.1375
14.7291 | -105.143
14.7328 | -105.1492
14.7365 | -105.1615
14.7401 | -105.1729
14.7438 | -105.1857
14.7475 | -105.1956
14.7511 | -105.2002
14.7548 | -105.2039
14.7585 | -105.2088
14.7622 | -105.2112
14.7658 | -105.2133
14.7695 | -105.2123
14.7732 | -105.2095
14.7769 | -105.2064
14.7805 | -105.2017
14.7842 | -105.196
14.7879 | -105.1842
14.7916 | -105.1726
14.7952 | -105.1608
14.7989 | -105.1429
14.8026 | -105.1247
14.8062 | -105.1025
14.8099 | -105.0769
14.8136 | -105.0554
14.8173 | -105.0341
14.8209 | -105.012
14.8246 | -104.9922
14.8283 | -104.9747
14.832 | -104.9558
14.8356 | -104.941
14.8393 | -104.9262
14.843 | -104.9124
14.8466 | -104.901
14.8503 | -104.8893
14.854 | -104.8777
14.8577 | -104.8688
14.8613 | -104.8619
14.865 | -104.8586
14.8687 | -104.8604
14.8724 | -104.8648
14.876 | -104.8701
14.8797 | -104.8785
14.8834 | -104.8898
14.8871 | -104.9045
14.8907 | -104.9198
14.8944 | -104.936
14.8981 | -104.9585
14.9017 | -104.9808
14.9054 | -105.0001
14.9091 | -105.0223
14.9128 | -105.0476
14.9164 | -105.0735
14.9201 | -105.0991
14.9238 | -105.1232
14.9275 | -105.144
14.9311 | -105.1647
14.9348 | -105.1848
14.9385 | -105.2059
14.9421 | -105.2263
14.9458 | -105.2444
14.9495 | -105.2565
14.9532 | -105.2668
14.9568 | -105.2739
14.9605 | -105.284
14.9642 | -105.2966
14.9679 | -105.3071
14.9715 | -105.3109
14.9752 | -105.3107
14.9789 | -105.3107
14.9826 | -105.3086
14.9862 | -105.3048
14.9899 | -105.2996
14.9936 | -105.2935
14.9972 | -105.2872
15.0009 | -105.2847
15.0046 | -105.2821
15.0083 | -105.2788
15.0119 | -105.2744
15.0156 | -105.2696
15.0193 | -105.2668
15.023 | -105.2634
15.0266 | -105.2555
15.0303 | -105.2476
15.034 | -105.242
15.0376 | -105.2345
15.0413 | -105.2295
15.045 | -105.2281
15.0487 | -105.2265
15.0523 | -105.2242
15.056 | -105.2232
15.0597 | -105.2192
15.0634 | -105.2133
15.067 | -105.212
15.0707 | -105.2111
15.0744 | -105.2067
15.0781 | -105.2008
15.0817 | -105.199
15.0854 | -105.1966
15.0891 | -105.1962
15.0927 | -105.1958
15.0964 | -105.1929
15.1001 | -105.1925
15.1038 | -105.1901
15.1074 | -105.1876
15.1111 | -105.1886
15.1148 | -105.1852
15.1185 | -105.1835
15.1221 | -105.1852
15.1258 | -105.1859
15.1295 | -105.1839
15.1332 | -105.1786
15.1368 | -105.1749
15.1405 | -105.1749
15.1442 | -105.1785
15.1478 | -105.1811
15.1515 | -105.1817
15.1552 | -105.1792
15.1589 | -105.1775
15.1625 | -105.1763
15.1662 | -105.1772
15.1699 | -105.1815
15.1736 | -105.1845
15.1772 | -105.1847
15.1809 | -105.1849
15.1846 | -105.186
15.1882 | -105.1906
15.1919 | -105.1957
15.1956 | -105.1973
15.1993 | -105.1982
15.2029 | -105.1995
15.2066 | -105.1977
15.2103 | -105.1938
15.214 | -105.1886
15.2176 | -105.1863
15.2213 | -105.1811
15.225 | -105.1762
15.2287 | -105.177
15.2323 | -105.176
15.236 | -105.1739
15.2397 | -105.1741
15.2433 | -105.1747
15.247 | -105.1733
15.2507 | -105.1722
15.2544 | -105.1728
15.258 | -105.1731
15.2617 | -105.1718
15.2654 | -105.1708
15.2691 | -105.1718
15.2727 | -105.1719
15.2764 | -105.1724
15.2801 | -105.1759
15.2837 | -105.1769
15.2874 | -105.1767
15.2911 | -105.1807
15.2948 | -105.1861
15.2984 | -105.187
15.3021 | -105.188
15.3058 | -105.1876
15.3095 | -105.1894
15.3131 | -105.1915
15.3168 | -105.19
15.3205 | -105.1924
15.3242 | -105.1987
15.3278 | -105.2021
15.3315 | -105.2051
15.3352 | -105.2063
15.3388 | -105.208
15.3425 | -105.2108
15.3462 | -105.2117
15.3499 | -105.2157
15.3535 | -105.2178
15.3572 | -105.2147
15.3609 | -105.2073
15.3646 | -105.1989
15.3682 | -105.1942
15.3719 | -105.1865
15.3756 | -105.1808
15.3792 | -105.1735
15.3829 | -105.1647
15.3866 | -105.1577
15.3903 | -105.1463
15.3939 | -105.1367
15.3976 | -105.1276
15.4013 | -105.1174
15.405 | -105.1087
15.4086 | -105.0997
15.4123 | -105.0917
15.416 | -105.0859
15.4197 | -105.0815
15.4233 | -105.0764
15.427 | -105.0752
15.4307 | -105.0764
15.4343 | -105.0774
15.438 | -105.0743
15.4417 | -105.0752
15.4454 | -105.0784
15.449 | -105.0794
15.4527 | -105.0828
15.4564 | -105.0878
15.4601 | -105.0928
15.4637 | -105.0997
15.4674 | -105.1072
15.4711 | -105.1156
15.4747 | -105.1238
15.4784 | -105.1314
15.4821 | -105.1439
15.4858 | -105.1569
15.4894 | -105.1684
15.4931 | -105.1811
15.4968 | -105.1933
15.5005 | -105.203
15.5041 | -105.2121
15.5078 | -105.2201
15.5115 | -104.8702
15.5152 | -103.899
15.5188 | -103.5322
15.5225 | -103.273
15.5262 | -102.994
15.5298 | -102.6967
15.5335 | -102.384
15.5372 | -102.0588
15.5409 | -101.7201
15.5445 | -101.3752
15.5482 | -101.0266
15.5519 | -100.6791
15.5556 | -100.3339
15.5592 | -99.9938
15.5629 | -99.6608
15.5666 | -99.3383
15.5702 | -99.0274
15.5739 | -98.7308
15.5776 | -98.4495
15.5813 | -98.185
15.5849 | -97.9376
15.5886 | -97.7085
15.5923 | -97.4995
15.596 | -97.3111
15.5996 | -97.1433
15.6033 | -96.9976
15.607 | -96.8735
15.6107 | -96.7713
15.6143 | -96.6922
15.618 | -96.6359
15.6217 | -96.6026
15.6253 | -96.5925
15.629 | -96.6051
15.6327 | -96.6407
15.6364 | -96.6993
15.64 | -96.7805
15.6437 | -96.8845
15.6474 | -97.0107
15.6511 | -97.1584
15.6547 | -97.3274
15.6584 | -97.5172
15.6621 | -97.7276
15.6657 | -97.9566
15.6694 | -98.204
15.6731 | -98.4684
15.6768 | -98.749
15.6804 | -99.0444
15.6841 | -99.3533
15.6878 | -99.6731
15.6915 | -100.0016
15.6951 | -100.3369
15.6988 | -100.6772
15.7025 | -101.0192
15.7062 | -101.3593
15.7098 | -101.6942
15.7135 | -102.0213
15.7172 | -102.3386
15.7208 | -102.6446
15.7245 | -102.9343
15.7282 | -103.2048
15.7319 | -103.4566
15.7355 | -103.9091
15.7392 | -104.8621
15.7429 | -105.072
15.7466 | -105.0766
15.7502 | -105.0781
15.7539 | -105.0801
15.7576 | -105.0832
15.7612 | -105.0844
15.7649 | -105.0866
15.7686 | -105.087
15.7723 | -105.0876
15.7759 | -105.0871
15.7796 | -105.0884
15.7833 | -105.0897
15.787 | -105.0909
15.7906 | -105.0962
15.7943 | -105.1006
15.798 | -105.1034
15.8017 | -105.1058
15.8053 | -105.1081
15.809 | -105.111
15.8127 | -105.1134
15.8163 | -105.115
15.82 | -105.1162
15.8237 | -105.1169
15.8274 | -105.1161
15.831 | -105.117
15.8347 | -105.1192
15.8384 | -105.1174
15.8421 | -105.1129
15.8457 | -105.1077
15.8494 | -105.1025
15.8531 | -105.0995
15.8567 | -105.0966
15.8604 | -105.0926
15.8641 | -105.0888
15.8678 | -105.0844
15.8714 | -105.0818
15.8751 | -105.0771
15.8788 | -105.0722
15.8825 | -105.067
15.8861 | -105.0611
15.8898 | -105.0554
15.8935 | -105.051
15.8972 | -105.0482
15.9008 | -105.0441
15.9045 | -105.0393
15.9082 | -105.0366
15.9118 | -105.0338
15.9155 | -105.0328
15.9192 | -105.0345
15.9229 | -105.0345
15.9265 | -105.0323
15.9302 | -105.0329
15.9339 | -105.0353
15.9376 | -105.0399
15.9412 | -105.0457
15.9449 | -105.0497
15.9486 | -105.051
15.9522 | -105.0549
15.9559 | -105.0613
15.9596 | -105.0663
15.9633 | -105.0727
15.9669 | -105.0818
15.9706 | -105.0873
15.9743 | -105.0931
15.978 | -105.1048
15.9816 | -105.1154
15.9853 | -105.126
15.989 | -105.1374
15.9927 | -105.152
15.9963 | -105.1651
16 | -105.1777
16.0037 | -105.1908
16.0073 | -105.2027
16.011 | -105.2138
16.0147 | -105.2255
16.0184 | -105.2344
16.022 | -105.2441
16.0257 | -105.2544
16.0294 | -105.2641
16.0331 | -105.2714
16.0367 | -105.2774
16.0404 | -105.2867
16.0441 | -105.2938
16.0478 | -105.2996
16.0514 | -105.3033
16.0551 | -105.3058
16.0588 | -105.3072
16.0624 | -105.3069
16.0661 | -105.3085
16.0698 | -105.3082
16.0735 | -105.3077
16.0771 | -105.3019
16.0808 | -105.2952
16.0845 | -105.2878
16.0882 | -105.2831
16.0918 | -105.2782
16.0955 | -105.2719
16.0992 | -105.2658
16.1028 | -105.262
16.1065 | -105.2576
16.1102 | -105.2544
16.1139 | -105.2492
16.1175 | -105.2455
16.1212 | -105.2397
16.1249 | -105.2334
16.1286 | -105.2277
16.1322 | -105.2249
16.1359 | -105.2186
16.1396 | -105.2096
16.1433 | -105.203
16.1469 | -105.1974
16.1506 | -105.1902
16.1543 | -105.1867
16.1579 | -105.1851
16.1616 | -105.181
16.1653 | -105.1766
16.169 | -105.1748
16.1726 | -105.1748
16.1763 | -105.1744
16.18 | -105.1764
16.1837 | -105.1767
16.1873 | -105.1794
16.191 | -105.1835
16.1947 | -105.184
16.1983 | -105.1852
16.202 | -105.1898
16.2057 | -105.1956
16.2094 | -105.1985
16.213 | -105.202
16.2167 | -105.2061
16.2204 | -105.2099
16.2241 | -105.2118
16.2277 | -105.2144
16.2314 | -105.2175
16.2351 | -105.218
16.2388 | -105.2182
16.2424 | -105.2211
16.2461 | -105.2255
16.2498 | -105.2314
16.2534 | -105.2366
16.2571 | -105.2414
16.2608 | -105.247
16.2645 | -105.2537
16.2681 | -105.2612
16.2718 | -105.2656
16.2755 | -105.271
16.2792 | -105.2785
16.2828 | -105.2869
16.2865 | -105.2947
16.2902 | -105.3036
16.2938 | -105.3143
16.2975 | -105.3238
16.3012 | -105.3361
16.3049 | -105.3481
16.3085 | -105.3621
16.3122 | -105.3734
16.3159 | -105.3848
16.3196 | -105.3946
16.3232 | -105.4037
16.3269 | -105.4117
16.3306 | -105.4196
16.3343 | -105.4292
16.3379 | -105.4384
16.3416 | -105.4451
16.3453 | -105.4531
16.3489 | -105.4618
16.3526 | -105.4688
16.3563 | -105.4752
16.36 | -105.4786
16.3636 | -105.4818
16.3673 | -105.4872
16.371 | -105.4914
16.3747 | -105.4927
16.3783 | -105.4924
16.382 | -105.494
16.3857 | -105.4931
16.3893 | -105.4896
16.393 | -105.4854
16.3967 | -105.4809
16.4004 | -105.4752
16.404 | -105.4661
16.4077 | -105.4556
16.4114 | -105.4446
16.4151 | -105.4366
16.4187 | -105.4255
16.4224 | -105.4137
16.4261 | -105.4007
16.4298 | -105.3869
16.4334 | -105.3736
16.4371 | -105.3618
16.4408 | -105.3488
16.4444 | -105.3356
16.4481 | -105.3236
16.4518 | -105.3108
16.4555 | -105.2996
16.4591 | -105.2896
16.4628 | -105.2777
16.4665 | -105.2649
16.4702 | -105.2546
16.4738 | -105.2461
16.4775 | -105.2358
16.4812 | -105.2241
16.4848 | -105.2127
16.4885 | -105.2037
16.4922 | -105.1962
16.4959 | -105.1891
16.4995 | -105.1817
16.5032 | -105.173
16.5069 | -105.1664
16.5106 | -105.1613
16.5142 | -105.1552
16.5179 | -105.1486
16.5216 | -105.1431
16.5253 | -105.1363
16.5289 | -105.1293
16.5326 | -105.1259
16.5363 | -105.1212
16.5399 | -105.1137
16.5436 | -105.1075
16.5473 | -105.1028
16.551 | -105.0982
16.5546 | -105.0924
16.5583 | -105.0852
16.562 | -105.0786
16.5657 | -105.0699
16.5693 | -105.0613
16.573 | -105.0526
16.5767 | -105.0466
16.5803 | -105.04
16.584 | -105.0341
16.5877 | -105.0293
16.5914 | -105.0252
16.595 | -105.0215
16.5987 | -105.0178
16.6024 | -105.0152
16.6061 | -105.0127
16.6097 | -105.0125
16.6134 | -105.0147
16.6171 | -105.018
16.6208 | -105.0228
16.6244 | -105.0303
16.6281 | -105.0381
16.6318 | -105.0457
16.6354 | -105.053
16.6391 | -105.0624
16.6428 | -105.0761
16.6465 | -105.0879
16.6501 | -105.0972
16.6538 | -105.1101
16.6575 | -105.1257
16.6612 | -105.1384
16.6648 | -105.1521
16.6685 | -105.1692
16.6722 | -105.1842
16.6758 | -105.1987
16.6795 | -105.2166
16.6832 | -105.2334
16.6869 | -105.2487
16.6905 | -105.264
16.6942 | -105.2793
16.6979 | -105.2935
16.7016 | -105.3104
16.7052 | -105.3258
16.7089 | -105.3386
16.7126 | -105.3513
16.7163 | -105.3623
16.7199 | -105.371
16.7236 | -105.3798
16.7273 | -105.3875
16.7309 | -105.3941
16.7346 | -105.3985
16.7383 | -105.4029
16.742 | -105.4075
16.7456 | -105.4087
16.7493 | -105.4075
16.753 | -105.4046
16.7567 | -105.4037
16.7603 | -105.3999
16.764 | -105.3966
16.7677 | -105.392
16.7714 | -105.383
16.775 | -105.3777
16.7787 | -105.3734
16.7824 | -105.3665
16.786 | -105.3595
16.7897 | -105.3514
16.7934 | -105.3433
16.7971 | -105.3359
16.8007 | -105.3296
16.8044 | -105.3241
16.8081 | -105.3172
16.8118 | -105.3111
16.8154 | -105.3047
16.8191 | -105.2981
16.8228 | -105.2906
16.8264 | -105.2846
16.8301 | -105.2806
16.8338 | -105.2767
16.8375 | -105.2738
16.8411 | -105.2709
16.8448 | -105.2705
16.8485 | -105.2702
16.8522 | -105.2675
16.8558 | -105.266
16.8595 | -105.2671
16.8632 | -105.2689
16.8668 | -105.2724
16.8705 | -105.2771
16.8742 | -105.284
16.8779 | -105.2896
16.8815 | -105.2898
16.8852 | -105.2946
16.8889 | -105.303
16.8926 | -105.3103
16.8962 | -105.3145
16.8999 | -105.3228
16.9036 | -105.333
16.9073 | -105.3404
16.9109 | -105.3477
16.9146 | -105.3538
16.9183 | -105.3571
16.9219 | -105.362
16.9256 | -105.3668
16.9293 | -105.3704
16.933 | -105.3719
16.9366 | -105.3741
16.9403 | -105.3746
16.944 | -105.3745
16.9477 | -105.3727
16.9513 | -105.3698
16.955 | -105.3656
16.9587 | -105.3638
16.9624 | -105.3624
16.966 | -105.3592
16.9697 | -105.3558
16.9734 | -105.3499
16.977 | -105.3428
16.9807 | -105.336
16.9844 | -105.3261
16.9881 | -105.3183
16.9917 | -105.3106
16.9954 | -105.3054
16.9991 | -105.2969
17.0028 | -105.2896
17.0064 | -105.2835
17.0101 | -105.2789
17.0138 | -105.2751
17.0174 | -105.268
17.0211 | -105.2629
17.0248 | -105.257
17.0285 | -105.2512
17.0321 | -105.2476
17.0358 | -105.2452
17.0395 | -105.2423
17.0432 | -105.239
17.0468 | -105.2348
17.0505 | -105.2313
17.0542 | -105.2286
17.0579 | -105.2259
17.0615 | -105.2243
17.0652 | -105.2214
17.0689 | -105.2207
17.0725 | -105.2202
17.0762 | -105.2196
17.0799 | -105.219
17.0836 | -105.2179
17.0872 | -105.2163
17.0909 | -105.2157
17.0946 | -105.2173
17.0983 | -105.216
17.1019 | -105.2132
17.1056 | -105.2137
17.1093 | -105.2167
17.1129 | -105.2203
17.1166 | -105.2219
17.1203 | -105.2264
17.124 | -105.2365
17.1276 | -105.2439
17.1313 | -105.2484
17.135 | -105.2536
17.1387 | -105.2599
17.1423 | -105.2673
17.146 | -105.2755
17.1497 | -105.2861
17.1534 | -105.2979
17.157 | -105.3108
17.1607 | -105.3242
17.1644 | -105.3363
17.168 | -105.3488
17.1717 | -105.3607
17.1754 | -105.3749
17.1791 | -105.3906
17.1827 | -105.4061
17.1864 | -105.4219
17.1901 | -105.4363
17.1938 | -105.4522
17.1974 | -105.4681
17.2011 | -105.4825
17.2048 | -105.499
17.2084 | -105.5159
17.2121 | -105.5308
17.2158 | -105.543
17.2195 | -105.5578
17.2231 | -105.5738
17.2268 | -105.589
17.2305 | -105.6061
17.2342 | -105.618
17.2378 | -105.6299
17.2415 | -105.6367
17.2452 | -105.6447
17.2489 | -105.6509
17.2525 | -105.6544
17.2562 | -105.6566
17.2599 | -105.6577
17.2635 | -105.6595
17.2672 | -105.6592
17.2709 | -105.6585
17.2746 | -105.6565
17.2782 | -105.6523
17.2819 | -105.6466
17.2856 | -105.6407
17.2893 | -105.6356
17.2929 | -105.6297
17.2966 | -105.6213
17.3003 | -105.61
17.3039 | -105.5975
17.3076 | -105.5808
17.3113 | -105.5677
17.315 | -105.5545
17.3186 | -105.5421
17.3223 | -105.5317
17.326 | -105.5181
17.3297 | -105.5054
17.3333 | -105.4927
17.337 | -105.4818
17.3407 | -105.4708
17.3444 | -105.4615
17.348 | -105.453
17.3517 | -105.4423
17.3554 | -105.4314
17.359 | -105.4218
17.3627 | -105.4143
17.3664 | -105.4057
17.3701 | -105.3953
17.3737 | -105.3876
17.3774 | -105.3826
17.3811 | -105.3763
17.3848 | -105.3717
17.3884 | -105.3658
17.3921 | -105.3622
17.3958 | -105.3604
17.3994 | -105.3592
17.4031 | -105.3585
17.4068 | -105.3576
17.4105 | -105.3574
17.4141 | -105.3578
17.4178 | -105.36
17.4215 | -105.3615
17.4252 | -105.36
17.4288 | -105.3588
17.4325 | -105.3561
17.4362 | -105.3541
17.4399 | -105.3515
17.4435 | -105.3476
17.4472 | -105.3446
17.4509 | -105.3406
17.4545 | -105.3349
17.4582 | -105.3311
17.4619 | -105.3279
17.4656 | -105.3243
17.4692 | -105.3204
17.4729 | -105.3144
17.4766 | -105.3058
17.4803 | -105.3002
17.4839 | -105.2989
17.4876 | -105.2949
17.4913 | -105.2949
17.4949 | -105.2926
17.4986 | -105.2889
17.5023 | -105.2892
17.506 | -105.2908
17.5096 | -105.2929
17.5133 | -105.2942
17.517 | -105.2949
17.5207 | -105.2981
17.5243 | -105.3006
17.528 | -105.3013
17.5317 | -105.3008
17.5354 | -105.3034
17.539 | -105.3065
17.5427 | -105.31
17.5464 | -105.3137
17.55 | -105.32
17.5537 | -105.3254
17.5574 | -105.3324
17.5611 | -105.3359
17.5647 | -105.3373
17.5684 | -105.3376
17.5721 | -105.3372
17.5758 | -105.3385
17.5794 | -105.3387
17.5831 | -105.3371
17.5868 | -105.3322
17.5905 | -105.3301
17.5941 | -105.3295
17.5978 | -105.3287
17.6015 | -105.3292
17.6051 | -105.3304
17.6088 | -105.3299
17.6125 | -105.3282
17.6162 | -105.3285
17.6198 | -105.3304
17.6235 | -105.3292
17.6272 | -105.3262
17.6309 | -105.3256
17.6345 | -105.3272
17.6382 | -105.3302
17.6419 | -105.3296
17.6455 | -105.3283
17.6492 | -105.3291
17.6529 | -105.3279
17.6566 | -105.3306
17.6602 | -105.3325
17.6639 | -105.3323
17.6676 | -105.3309
17.6713 | -105.3306
17.6749 | -105.3294
17.6786 | -105.3278
17.6823 | -105.3274
17.6859 | -105.3273
17.6896 | -105.3243
17.6933 | -105.32
17.697 | -105.3147
17.7006 | -105.3108
17.7043 | -105.3054
17.708 | -105.2976
17.7117 | -105.2895
17.7153 | -105.2809
17.719 | -105.2712
17.7227 | -105.2619
17.7264 | -105.2553
17.73 | -105.2477
17.7337 | -105.2404
17.7374 | -105.233
17.741 | -105.2233
17.7447 | -105.2126
17.7484 | -105.2055
17.7521 | -105.202
17.7557 | -105.1952
17.7594 | -105.1886
17.7631 | -105.1832
17.7668 | -105.1789
17.7704 | -105.1763
17.7741 | -105.1737
17.7778 | -105.1716
17.7815 | -105.1717
17.7851 | -105.1748
17.7888 | -105.1751
17.7925 | -105.1751
17.7961 | -105.1761
17.7998 | -105.1801
17.8035 | -105.1859
17.8072 | -105.1895
17.8108 | -105.1929
17.8145 | -105.1972
17.8182 | -105.2032
17.8219 | -105.2108
17.8255 | -105.219
17.8292 | -105.2294
17.8329 | -105.2385
17.8365 | -105.2465
17.8402 | -105.2551
17.8439 | -105.2629
17.8476 | -105.2713
17.8512 | -105.2792
17.8549 | -105.2844
17.8586 | -105.2893
17.8623 | -105.2955
17.8659 | -105.3027
17.8696 | -105.3068
17.8733 | -105.3041
17.877 | -105.3012
17.8806 | -105.3025
17.8843 | -105.3023
17.888 | -105.2964
17.8916 | -105.2908
17.8953 | -105.2862
17.899 | -105.279
17.9027 | -105.27
17.9063 | -105.2623
17.91 | -105.2532
17.9137 | -105.2442
17.9174 | -105.2366
17.921 | -105.2285
17.9247 | -105.2175
17.9284 | -105.2076
17.932 | -105.2001
17.9357 | -105.1937
17.9394 | -105.1868
17.9431 | -105.18
17.9467 | -105.1749
17.9504 | -105.1726
17.9541 | -105.1707
17.9578 | -105.1701
17.9614 | -105.1709
17.9651 | -105.1726
17.9688 | -105.1748
17.9725 | -105.1793
17.9761 | -105.1864
17.9798 | -105.1942
17.9835 | -105.2023
17.9871 | -105.2093
17.9908 | -105.2185
17.9945 | -105.2316
17.9982 | -105.2461
18.0018 | -105.26
18.0055 | -105.2722
18.0092 | -105.2847
18.0129 | -105.3011
18.0165 | -105.3195
18.0202 | -105.3342
18.0239 | -105.3497
18.0275 | -105.3649
18.0312 | -105.3796
18.0349 | -105.3933
18.0386 | -105.4064
18.0422 | -105.4171
18.0459 | -105.4287
18.0496 | -105.4394
18.0533 | -105.4478
18.0569 | -105.4526
18.0606 | -105.4565
18.0643 | -105.4653
18.068 | -105.4733
18.0716 | -105.477
18.0753 | -105.4773
18.079 | -105.4758
18.0826 | -105.4727
18.0863 | -105.4677
18.09 | -105.4644
18.0937 | -105.4605
18.0973 | -105.4556
18.101 | -105.4457
18.1047 | -105.4304
18.1084 | -105.415
18.112 | -105.4017
18.1157 | -105.3863
18.1194 | -105.3701
18.123 | -105.3523
18.1267 | -105.336
18.1304 | -105.3234
18.1341 | -105.3086
18.1377 | -105.2855
18.1414 | -105.2589
18.1451 | -105.2321
18.1488 | -105.2034
18.1524 | -105.1732
18.1561 | -105.1408
18.1598 | -105.1071
18.1635 | -105.0719
18.1671 | -105.0369
18.1708 | -105.0007
18.1745 | -104.9645
18.1781 | -104.927
18.1818 | -104.8892
18.1855 | -104.8507
18.1892 | -104.8123
18.1928 | -104.7715
18.1965 | -104.7327
18.2002 | -104.6975
18.2039 | -104.6625
18.2075 | -104.6263
18.2112 | -104.5892
18.2149 | -104.5567
18.2185 | -104.5263
18.2222 | -104.4973
18.2259 | -104.4711
18.2296 | -104.4476
18.2332 | -104.4244
18.2369 | -104.4045
18.2406 | -104.3888
18.2443 | -104.3772
18.2479 | -104.3667
18.2516 | -104.3562
18.2553 | -104.3518
18.259 | -104.3502
18.2626 | -104.3535
18.2663 | -104.3601
18.27 | -104.3661
18.2736 | -104.3752
18.2773 | -104.3885
18.281 | -104.4064
18.2847 | -104.4262
18.2883 | -104.4487
18.292 | -104.4733
18.2957 | -104.4994
18.2994 | -104.5265
18.303 | -104.5557
18.3067 | -104.5843
18.3104 | -104.6143
18.3141 | -104.6468
18.3177 | -104.6795
18.3214 | -104.7114
18.3251 | -104.7418
18.3287 | -104.7726
18.3324 | -104.8052
18.3361 | -104.8372
18.3398 | -104.8671
18.3434 | -104.8956
18.3471 | -104.9226
18.3508 | -104.9476
18.3545 | -104.9713
18.3581 | -104.9898
18.3618 | -105.0078
18.3655 | -105.0248
18.3691 | -105.0416
18.3728 | -105.0546
18.3765 | -105.0655
18.3802 | -105.0721
18.3838 | -105.0778
18.3875 | -105.0787
18.3912 | -105.0755
18.3949 | -105.0711
18.3985 | -105.0675
18.4022 | -105.0636
18.4059 | -105.0579
18.4095 | -105.0478
18.4132 | -105.0365
18.4169 | -105.0236
18.4206 | -105.0096
18.4242 | -104.9952
18.4279 | -104.978
18.4316 | -104.9598
18.4353 | -104.9402
18.4389 | -104.9235
18.4426 | -104.9078
18.4463 | -104.8929
18.45 | -104.8774
18.4536 | -104.8665
18.4573 | -104.856
18.461 | -104.8432
18.4646 | -104.8326
18.4683 | -104.8245
18.472 | -104.8212
18.4757 | -104.8192
18.4793 | -104.8181
18.483 | -104.8187
18.4867 | -104.8221
18.4904 | -104.8257
18.494 | -104.8334
18.4977 | -104.842
18.5014 | -104.8538
18.5051 | -104.8676
18.5087 | -104.8824
18.5124 | -104.8993
18.5161 | -104.914
18.5197 | -104.9295
18.5234 | -104.946
18.5271 | -104.9614
18.5308 | -104.9785
18.5344 | -104.9964
18.5381 | -105.0135
18.5418 | -105.0306
18.5455 | -105.0437
18.5491 | -105.0585
18.5528 | -105.0735
18.5565 | -105.0862
18.5601 | -105.1015
18.5638 | -105.117
18.5675 | -105.1285
18.5712 | -105.1396
18.5748 | -105.1498
18.5785 | -105.1586
18.5822 | -105.1652
18.5859 | -105.1699
18.5895 | -105.1722
18.5932 | -105.1729
18.5969 | -105.1709
18.6006 | -105.1672
18.6042 | -105.1623
18.6079 | -105.1564
18.6116 | -105.1511
18.6152 | -105.1432
18.6189 | -105.1332
18.6226 | -105.1264
18.6263 | -105.1191
18.6299 | -105.1085
18.6336 | -105.092
18.6373 | -105.0713
18.641 | -105.0531
18.6446 | -105.0379
18.6483 | -105.02
18.652 | -105.0021
18.6556 | -104.985
18.6593 | -104.9661
18.663 | -104.9469
18.6667 | -104.9279
18.6703 | -104.9065
18.674 | -104.8867
18.6777 | -104.8653
18.6814 | -104.8459
18.685 | -104.8261
18.6887 | -104.8064
18.6924 | -104.7855
18.6961 | -104.7674
18.6997 | -104.7501
18.7034 | -104.7331
18.7071 | -104.7176
18.7107 | -104.7036
18.7144 | -104.6908
18.7181 | -104.6778
18.7218 | -104.6672
18.7254 | -104.6598
18.7291 | -104.6542
18.7328 | -104.6476
18.7365 | -104.6397
18.7401 | -104.6328
18.7438 | -104.6281
18.7475 | -104.6268
18.7511 | -104.6266
18.7548 | -104.6297
18.7585 | -104.6332
18.7622 | -104.6365
18.7658 | -104.6404
18.7695 | -104.6476
18.7732 | -104.6576
18.7769 | -104.6687
18.7805 | -104.6803
18.7842 | -104.6919
18.7879 | -104.705
18.7916 | -104.7213
18.7952 | -104.7388
18.7989 | -104.7557
18.8026 | -104.7732
18.8062 | -104.7902
18.8099 | -104.808
18.8136 | -104.8274
18.8173 | -104.8467
18.8209 | -104.8637
18.8246 | -104.881
18.8283 | -104.8998
18.832 | -104.9184
18.8356 | -104.9395
18.8393 | -104.959
18.843 | -104.9757
18.8466 | -104.9891
18.8503 | -105.0039
18.854 | -105.0187
18.8577 | -105.031
18.8613 | -105.0425
18.865 | -105.0513
18.8687 | -105.0583
18.8724 | -105.0662
18.876 | -105.0716
18.8797 | -105.0722
18.8834 | -105.0736
18.8871 | -105.0745
18.8907 | -105.0726
18.8944 | -105.0663
18.8981 | -105.0614
18.9017 | -105.0559
18.9054 | -105.0519
18.9091 | -105.0451
18.9128 | -105.0359
18.9164 | -105.0256
18.9201 | -105.0143
18.9238 | -105.0017
18.9275 | -104.9883
18.9311 | -104.9757
18.9348 | -104.9616
18.9385 | -104.9493
18.9421 | -104.9382
18.9458 | -104.9297
18.9495 | -104.919
18.9532 | -104.9083
18.9568 | -104.8977
18.9605 | -104.8897
18.9642 | -104.8809
18.9679 | -104.8729
18.9715 | -104.8645
18.9752 | -104.8551
18.9789 | -104.8475
18.9826 | -104.8408
18.9862 | -104.8367
18.9899 | -104.8348
18.9936 | -104.8334
18.9972 | -104.8329
19.0009 | -104.8376
19.0046 | -104.8455
19.0083 | -104.8538
19.0119 | -104.8633
19.0156 | -104.872
19.0193 | -104.8817
19.023 | -104.8905
19.0266 | -104.9007
19.0303 | -104.9144
19.034 | -104.9268
19.0376 | -104.9391
19.0413 | -104.9545
19.045 | -104.9722
19.0487 | -104.9899
19.0523 | -105.0084
19.056 | -105.0257
19.0597 | -105.0431
19.0634 | -105.0605
19.067 | -105.0771
19.0707 | -105.0948
19.0744 | -105.1127
19.0781 | -105.1295
19.0817 | -105.1463
19.0854 | -105.1622
19.0891 | -105.1768
19.0927 | -105.1941
19.0964 | -105.2104
19.1001 | -105.2215
19.1038 | -105.2326
19.1074 | -105.2429
19.1111 | -105.2532
19.1148 | -105.2614
19.1185 | -105.2707
19.1221 | -105.2795
19.1258 | -105.2847
19.1295 | -105.2911
19.1332 | -105.2948
19.1368 | -105.2995
19.1405 | -105.3019
19.1442 | -105.3008
19.1478 | -105.302
19.1515 | -105.3037
19.1552 | -105.3024
19.1589 | -105.2977
19.1625 | -105.292
19.1662 | -105.2888
19.1699 | -105.2861
19.1736 | -105.2816
19.1772 | -105.2751
19.1809 | -105.266
19.1846 | -105.2577
19.1882 | -105.2509
19.1919 | -105.2444
19.1956 | -105.2387
19.1993 | -105.2333
19.2029 | -105.2261
19.2066 | -105.2214
19.2103 | -105.2185
19.214 | -105.2131
19.2176 | -105.207
19.2213 | -105.2048
19.225 | -105.2017
19.2286 | -105.1999
19.2323 | -105.198
19.236 | -105.1953
19.2397 | -105.1949
19.2433 | -105.1969
19.247 | -105.1997
19.2507 | -105.2015
19.2544 | -105.2024
19.258 | -105.2025
19.2617 | -105.203
19.2654 | -105.2069
19.2691 | -105.2122
19.2727 | -105.2152
19.2764 | -105.2163
19.2801 | -105.2195
19.2837 | -105.2249
19.2874 | -105.2316
19.2911 | -105.2389
19.2948 | -105.2476
19.2984 | -105.2579
19.3021 | -105.2677
19.3058 | -105.2767
19.3095 | -105.2828
19.3131 | -105.2883
19.3168 | -105.2959
19.3205 | -105.3042
19.3242 | -105.3133
19.3278 | -105.3207
19.3315 | -105.3278
19.3352 | -105.3345
19.3388 | -105.3419
19.3425 | -105.3486
19.3462 | -105.3541
19.3499 | -105.359
19.3535 | -105.3672
19.3572 | -105.3724
19.3609 | -105.3761
19.3646 | -105.3786
19.3682 | -105.3814
19.3719 | -105.3849
19.3756 | -105.3858
19.3792 | -105.3859
19.3829 | -105.3861
19.3866 | -105.3842
19.3903 | -105.3824
19.3939 | -105.381
19.3976 | -105.3763
19.4013 | -105.372
19.405 | -105.3672
19.4086 | -105.3625
19.4123 | -105.3559
19.416 | -105.3506
19.4197 | -105.345
19.4233 | -105.3359
19.427 | -105.3245
19.4307 | -105.3137
19.4343 | -105.3022
19.438 | -105.2879
19.4417 | -105.2751
19.4454 | -105.2611
19.449 | -105.2457
19.4527 | -105.2316
19.4564 | -105.2178
19.4601 | -105.2007
19.4637 | -105.1852
19.4674 | -105.1698
19.4711 | -105.1531
19.4747 | -105.1375
19.4784 | -105.1224
19.4821 | -105.1056
19.4858 | -105.0898
19.4894 | -105.0739
19.4931 | -105.0592
19.4968 | -105.0439
19.5005 | -105.0318
19.5041 | -105.0207
19.5078 | -105.0108
19.5115 | -105.0047
19.5152 | -104.9973
19.5188 | -104.989
19.5225 | -104.9842
19.5262 | -104.9824
19.5298 | -104.9816
19.5335 | -104.9831
19.5372 | -104.9833
19.5409 | -104.9849
19.5445 | -104.987
19.5482 | -104.9928
19.5519 | -105.0009
19.5556 | -105.0081
19.5592 | -105.0164
19.5629 | -105.0267
19.5666 | -105.0378
19.5702 | -105.0491
19.5739 | -105.059
19.5776 | -105.07
19.5813 | -105.0838
19.5849 | -105.0985
19.5886 | -105.1119
19.5923 | -105.1224
19.596 | -105.1305
19.5996 | -105.1403
19.6033 | -105.1465
19.607 | -105.1582
19.6107 | -105.1682
19.6143 | -105.1762
19.618 | -105.1842
19.6217 | -105.191
19.6253 | -105.1968
19.629 | -105.2027
19.6327 | -105.2058
19.6364 | -105.2089
19.64 | -105.2113
19.6437 | -105.2124
19.6474 | -105.2117
19.6511 | -105.2086
19.6547 | -105.2068
19.6584 | -105.203
19.6621 | -105.1995
19.6657 | -105.1967
19.6694 | -105.1943
19.6731 | -105.188
19.6768 | -105.1816
19.6804 | -105.174
19.6841 | -105.1665
19.6878 | -105.1585
19.6915 | -105.1513
19.6951 | -105.1455
19.6988 | -105.1404
19.7025 | -105.1354
19.7062 | -105.1322
19.7098 | -105.1309
19.7135 | -105.1326
19.7172 | -105.1336
19.7208 | -105.1325
19.7245 | -105.1311
19.7282 | -105.1333
19.7319 | -105.1362
19.7355 | -105.1391
19.7392 | -105.1444
19.7429 | -105.1491
19.7466 | -105.1529
19.7502 | -105.1591
19.7539 | -105.1661
19.7576 | -105.171
19.7612 | -105.1793
19.7649 | -105.1871
19.7686 | -105.1958
19.7723 | -105.2061
19.7759 | -105.2149
19.7796 | -105.2219
19.7833 | -105.2288
19.787 | -105.236
19.7906 | -105.2414
19.7943 | -105.249
19.798 | -105.2572
19.8017 | -105.2616
19.8053 | -105.2654
19.809 | -105.2716
19.8127 | -105.2745
19.8163 | -105.2784
19.82 | -105.2793
19.8237 | -105.2796
19.8274 | -105.2772
19.831 | -105.2749
19.8347 | -105.274
19.8384 | -105.2713
19.8421 | -105.2665
19.8457 | -105.2619
19.8494 | -105.2585
19.8531 | -105.2568
19.8567 | -105.2508
19.8604 | -105.241
19.8641 | -105.2335
19.8678 | -105.2261
19.8714 | -105.2185
19.8751 | -105.2134
19.8788 | -105.2047
19.8825 | -105.1949
19.8861 | -105.1835
19.8898 | -105.1714
19.8935 | -105.1598
19.8972 | -105.1483
19.9008 | -105.1352
19.9045 | -105.1232
19.9082 | -105.1123
19.9118 | -105.1008
19.9155 | -105.0899
19.9192 | -105.0793
19.9229 | -105.0709
19.9265 | -105.0643
19.9302 | -105.0556
19.9339 | -105.0475
19.9376 | -105.0391
19.9412 | -105.0319
19.9449 | -105.0224
19.9486 | -105.0132
19.9522 | -105.0046
19.9559 | -104.9979
19.9596 | -104.9894
19.9633 | -104.9822
19.9669 | -104.9761
19.9706 | -104.9707
19.9743 | -104.9639
19.978 | -104.9535
19.9816 | -104.9456
19.9853 | -104.9398
19.989 | -104.934
19.9927 | -104.9283
19.9963 | -104.9235
20 | -104.9194
20.0037 | -104.9142
20.0073 | -104.9099
20.011 | -104.9051
20.0147 | -104.8981
20.0184 | -104.8909
20.022 | -104.8839
20.0257 | -104.877
20.0294 | -104.8708
20.0331 | -104.8658
20.0367 | -104.8589
20.0404 | -104.8523
20.0441 | -104.8439
20.0478 | -104.8367
20.0514 | -104.8291
20.0551 | -104.8208
20.0588 | -104.814
20.0624 | -104.8082
20.0661 | -104.8011
20.0698 | -104.7948
20.0735 | -104.7904
20.0771 | -104.7861
20.0808 | -104.7814
20.0845 | -104.7787
20.0882 | -104.7778
20.0918 | -104.7759
20.0955 | -104.7777
20.0992 | -104.7791
20.1028 | -104.7825
20.1065 | -104.7883
20.1102 | -104.7952
20.1139 | -104.802
20.1175 | -104.8072
20.1212 | -104.8116
20.1249 | -104.8182
20.1286 | -104.8257
20.1322 | -104.8346
20.1359 | -104.8469
20.1396 | -104.8563
20.1433 | -104.8683
20.1469 | -104.8806
20.1506 | -104.8906
20.1543 | -104.8998
20.1579 | -104.9086
20.1616 | -104.9174
20.1653 | -104.9262
20.169 | -104.9365
20.1726 | -104.9468
20.1763 | -104.957
20.18 | -104.9666
20.1837 | -104.9762
20.1873 | -104.9853
20.191 | -104.9926
20.1947 | -104.9981
20.1983 | -105.0041
20.202 | -105.0088
20.2057 | -105.0138
20.2094 | -105.0186
20.213 | -105.0221
20.2167 | -105.0253
20.2204 | -105.0304
20.2241 | -105.0357
20.2277 | -105.04
20.2314 | -105.045
20.2351 | -105.051
20.2388 | -105.0546
20.2424 | -105.0555
20.2461 | -105.0586
20.2498 | -105.0629
20.2534 | -105.0668
20.2571 | -105.0685
20.2608 | -105.0701
20.2645 | -105.0723
20.2681 | -105.071
20.2718 | -105.0709
20.2755 | -105.0698
20.2792 | -105.0682
20.2828 | -105.0682
20.2865 | -105.0693
20.2902 | -105.0714
20.2938 | -105.0728
20.2975 | -105.0745
20.3012 | -105.0786
20.3049 | -105.0822
20.3085 | -105.0848
20.3122 | -105.0883
20.3159 | -105.0937
20.3196 | -105.0962
20.3232 | -105.1009
20.3269 | -105.1068
20.3306 | -105.1122
20.3343 | -105.1169
20.3379 | -105.1244
20.3416 | -105.1284
20.3453 | -105.1338
20.3489 | -105.1406
20.3526 | -105.1477
20.3563 | -105.1523
20.36 | -105.1568
20.3636 | -105.1609
20.3673 | -105.1666
20.371 | -105.1741
20.3747 | -105.1798
20.3783 | -105.1841
20.382 | -105.1877
20.3857 | -105.1938
20.3893 | -105.1995
20.393 | -105.2034
20.3967 | -105.2068
20.4004 | -105.2107
20.404 | -105.2184
20.4077 | -105.2279
20.4114 | -105.2353
20.4151 | -105.2408
20.4187 | -105.2448
20.4224 | -105.251
20.4261 | -105.2557
20.4298 | -105.2581
20.4334 | -105.2594
20.4371 | -105.2628
20.4408 | -105.2666
20.4444 | -105.2691
20.4481 | -105.2712
20.4518 | -105.2721
20.4555 | -105.2742
20.4591 | -105.2735
20.4628 | -105.2723
20.4665 | -105.2689
20.4702 | -105.2658
20.4738 | -105.2615
20.4775 | -105.2582
20.4812 | -105.2544
20.4848 | -105.2482
20.4885 | -105.2435
20.4922 | -105.2391
20.4959 | -105.2333
20.4995 | -105.2256
20.5032 | -105.2179
20.5069 | -105.2086
20.5106 | -105.198
20.5142 | -105.1882
20.5179 | -105.1776
20.5216 | -105.1658
20.5253 | -105.1532
20.5289 | -105.1425
20.5326 | -105.1325
20.5363 | -105.1228
20.5399 | -105.1118
20.5436 | -105.0987
20.5473 | -105.0876
20.551 | -105.077
20.5546 | -105.0686
20.5583 | -105.0592
20.562 | -105.0474
20.5657 | -105.0359
20.5693 | -105.0235
20.573 | -105.0075
20.5767 | -104.9934
20.5803 | -104.9807
20.584 | -104.9681
20.5877 | -104.955
20.5914 | -104.9415
20.595 | -104.9305
20.5987 | -104.9182
20.6024 | -104.9067
20.6061 | -104.8936
20.6097 | -104.8828
20.6134 | -104.8726
20.6171 | -104.8621
20.6208 | -104.8533
20.6244 | -104.844
20.6281 | -104.8344
20.6318 | -104.8279
20.6354 | -104.8237
20.6391 | -104.8185
20.6428 | -104.8132
20.6465 | -104.8117
20.6501 | -104.811
20.6538 | -104.8118
20.6575 | -104.8131
20.6612 | -104.8148
20.6648 | -104.8249
20.6685 | -104.827
20.6722 | -104.8316
20.6758 | -104.838
20.6795 | -104.8434
20.6832 | -104.849
20.6869 | -104.8575
20.6905 | -104.867
20.6942 | -104.8762
20.6979 | -104.883
20.7016 | -104.8913
20.7052 | -104.9018
20.7089 | -104.9134
20.7126 | -104.9242
20.7163 | -104.9341
20.7199 | -104.946
20.7236 | -104.9578
20.7273 | -104.9715
20.7309 | -104.9836
20.7346 | -104.9921
20.7383 | -105.0009
20.742 | -105.0113
20.7456 | -105.0219
20.7493 | -105.0323
20.753 | -105.0419
20.7567 | -105.0506
20.7603 | -105.0609
20.764 | -105.0688
20.7677 | -105.078
20.7714 | -105.0862
20.775 | -105.0918
20.7787 | -105.0985
20.7824 | -105.1028
20.786 | -105.108
20.7897 | -105.1126
20.7934 | -105.1164
20.7971 | -105.1187
20.8007 | -105.1244
20.8044 | -105.129
20.8081 | -105.1311
20.8118 | -105.1348
20.8154 | -105.1384
20.8191 | -105.1405
20.8228 | -105.1388
20.8264 | -105.139
20.8301 | -105.1385
20.8338 | -105.1351
20.8375 | -105.1318
20.8411 | -105.1302
20.8448 | -105.129
20.8485 | -105.1233
20.8522 | -105.1139
20.8558 | -105.1055
20.8595 | -105.0998
20.8632 | -105.0927
20.8668 | -105.086
20.8705 | -105.0768
20.8742 | -105.0679
20.8779 | -105.0594
20.8815 | -105.0488
20.8852 | -105.0374
20.8889 | -105.0256
20.8926 | -105.0158
20.8962 | -105.0033
20.8999 | -104.992
20.9036 | -104.9801
20.9073 | -104.968
20.9109 | -104.957
20.9146 | -104.9456
20.9183 | -104.9336
20.9219 | -104.9215
20.9256 | -104.9107
20.9293 | -104.897
20.933 | -104.8857
20.9366 | -104.8762
20.9403 | -104.8639
20.944 | -104.8519
20.9477 | -104.8407
20.9513 | -104.8309
20.955 | -104.8233
20.9587 | -104.8185
20.9624 | -104.8142
20.966 | -104.8102
20.9697 | -104.8087
20.9734 | -104.807
20.977 | -104.8073
20.9807 | -104.8107
20.9844 | -104.8086
20.9881 | -104.8108
20.9917 | -104.8154
20.9954 | -104.8224
20.9991 | -104.8284
21.0028 | -104.8355
21.0064 | -104.8426
21.0101 | -104.8498
21.0138 | -104.8582
21.0174 | -104.8686
21.0211 | -104.8789
21.0248 | -104.8881
21.0285 | -104.8988
21.0321 | -104.9123
21.0358 | -104.9259
21.0395 | -104.9382
21.0432 | -104.9502
21.0468 | -104.9612
21.0505 | -104.9738
21.0542 | -104.9853
21.0579 | -104.9958
21.0615 | -105.0077
21.0652 | -105.0175
21.0689 | -105.0265
21.0725 | -105.0375
21.0762 | -105.049
21.0799 | -105.0593
21.0836 | -105.0675
21.0872 | -105.0766
21.0909 | -105.0861
21.0946 | -105.0943
21.0983 | -105.0982
21.1019 | -105.1021
21.1056 | -105.1068
21.1093 | -105.1096
21.1129 | -105.1122
21.1166 | -105.117
21.1203 | -105.123
21.124 | -105.1256
21.1276 | -105.128
21.1313 | -105.1326
21.135 | -105.1365
21.1387 | -105.1373
21.1423 | -105.1374
21.146 | -105.1393
21.1497 | -105.1445
21.1534 | -105.1482
21.157 | -105.1497
21.1607 | -105.1496
21.1644 | -105.1495
21.168 | -105.1505
21.1717 | -105.1493
21.1754 | -105.1455
21.1791 | -105.1452
21.1827 | -105.1455
21.1864 | -105.1445
21.1901 | -105.144
21.1938 | -105.1443
21.1974 | -105.1438
21.2011 | -105.1439
21.2048 | -105.144
21.2084 | -105.143
21.2121 | -105.1423
21.2158 | -105.1417
21.2195 | -105.1422
21.2231 | -105.1413
21.2268 | -105.1396
21.2305 | -105.1404
21.2342 | -105.139
21.2378 | -105.1378
21.2415 | -105.1391
21.2452 | -105.1413
21.2489 | -105.1418
21.2525 | -105.143
21.2562 | -105.1432
21.2599 | -105.1421
21.2635 | -105.1397
21.2672 | -105.1401
21.2709 | -105.1366
21.2746 | -105.1342
21.2782 | -105.134
21.2819 | -105.1309
21.2856 | -105.1316
21.2893 | -105.1309
21.2929 | -105.1287
21.2966 | -105.1251
21.3003 | -105.1223
21.3039 | -105.1209
21.3076 | -105.1204
21.3113 | -105.1211
21.315 | -105.121
21.3186 | -105.1182
21.3223 | -105.1169
21.326 | -105.1168
21.3297 | -105.1144
21.3333 | -105.1108
21.337 | -105.1088
21.3407 | -105.1065
21.3444 | -105.1047
21.348 | -105.1043
21.3517 | -105.1033
21.3554 | -105.1035
21.359 | -105.1038
21.3627 | -105.1043
21.3664 | -105.1027
21.3701 | -105.1027
21.3737 | -105.103
21.3774 | -105.1027
21.3811 | -105.1041
21.3848 | -105.1069
21.3884 | -105.11
21.3921 | -105.1118
21.3958 | -105.1155
21.3994 | -105.1206
21.4031 | -105.1259
21.4068 | -105.131
21.4105 | -105.1354
21.4141 | -105.1389
21.4178 | -105.1419
21.4215 | -105.1453
21.4252 | -105.1486
21.4288 | -105.1509
21.4325 | -105.154
21.4362 | -105.1592
21.4399 | -105.1644
21.4435 | -105.171
21.4472 | -105.1784
21.4509 | -105.1861
21.4545 | -105.1924
21.4582 | -105.1954
21.4619 | -105.2009
21.4656 | -105.206
21.4692 | -105.2074
21.4729 | -105.21
21.4766 | -105.218
21.4803 | -105.2245
21.4839 | -105.2271
21.4876 | -105.2327
21.4913 | -105.2372
21.4949 | -105.2404
21.4986 | -105.2439
21.5023 | -105.2484
21.506 | -105.2508
21.5096 | -105.2522
21.5133 | -105.2528
21.517 | -105.2537
21.5207 | -105.2534
21.5243 | -105.252
21.528 | -105.2522
21.5317 | -105.2513
21.5354 | -105.2478
21.539 | -105.2445
21.5427 | -105.2419
21.5464 | -105.2381
21.55 | -105.2308
21.5537 | -105.2235
21.5574 | -105.2149
21.5611 | -105.2047
21.5647 | -105.194
21.5684 | -105.1834
21.5721 | -105.1704
21.5758 | -105.1607
21.5794 | -105.1484
21.5831 | -105.1324
21.5868 | -105.119
21.5905 | -105.1053
21.5941 | -105.0885
21.5978 | -105.0729
21.6015 | -105.0572
21.6051 | -105.0412
21.6088 | -105.0232
21.6125 | -105.0056
21.6162 | -104.9898
21.6198 | -104.9742
21.6235 | -104.9591
21.6272 | -104.9437
21.6309 | -104.9298
21.6345 | -104.9175
21.6382 | -104.906
21.6419 | -104.8932
21.6455 | -104.8807
21.6492 | -104.869
21.6529 | -104.8581
21.6566 | -104.849
21.6602 | -104.8422
21.6639 | -104.8357
21.6676 | -104.8314
21.6713 | -104.8285
21.6749 | -104.825
21.6786 | -104.8223
21.6823 | -104.8222
21.6859 | -104.8219
21.6896 | -104.8211
21.6933 | -104.8217
21.697 | -104.825
21.7006 | -104.8298
21.7043 | -104.8359
21.708 | -104.8414
21.7117 | -104.8466
21.7153 | -104.8536
21.719 | -104.8629
21.7227 | -104.8712
21.7264 | -104.8823
21.73 | -104.8906
21.7337 | -104.8985
21.7374 | -104.9055
21.741 | -104.9144
21.7447 | -104.9225
21.7484 | -104.9326
21.7521 | -104.9422
21.7557 | -104.951
21.7594 | -104.9583
21.7631 | -104.9643
21.7668 | -104.9721
21.7704 | -104.981
21.7741 | -104.989
21.7778 | -104.9973
21.7815 | -105.0053
21.7851 | -105.0086
21.7888 | -105.0103
21.7925 | -105.0129
21.7961 | -105.0137
21.7998 | -105.0149
21.8035 | -105.0158
21.8072 | -105.0146
21.8108 | -105.0124
21.8145 | -105.0121
21.8182 | -105.0125
21.8219 | -105.0133
21.8255 | -105.0114
21.8292 | -105.008
21.8329 | -105.003
21.8365 | -104.9968
21.8402 | -104.9904
21.8439 | -104.981
21.8476 | -104.9709
21.8512 | -104.9612
21.8549 | -104.9505
21.8586 | -104.943
21.8623 | -104.9342
21.8659 | -104.9235
21.8696 | -104.9127
21.8733 | -104.9033
21.877 | -104.8916
21.8806 | -104.8795
21.8843 | -104.8655
21.888 | -104.8532
21.8916 | -104.8406
21.8953 | -104.827
21.899 | -104.8126
21.9027 | -104.7984
21.9063 | -104.7845
21.91 | -104.7712
21.9137 | -104.7583
21.9174 | -104.7448
21.921 | -104.7324
21.9247 | -104.716
21.9284 | -104.7005
21.932 | -104.6877
21.9357 | -104.6758
21.9394 | -104.6643
21.9431 | -104.6538
21.9467 | -104.6444
21.9504 | -104.6353
21.9541 | -104.627
21.9578 | -104.6176
21.9614 | -104.6086
21.9651 | -104.5998
21.9688 | -104.5917
21.9725 | -104.5857
21.9761 | -104.5791
21.9798 | -104.5744
21.9835 | -104.571
21.9871 | -104.5701
21.9908 | -104.5682
21.9945 | -104.5678
21.9982 | -104.5696
22.0018 | -104.572
22.0055 | -104.575
22.0092 | -104.5792
22.0129 | -104.5831
22.0165 | -104.5883
22.0202 | -104.5961
22.0239 | -104.6048
22.0275 | -104.6121
22.0312 | -104.6196
22.0349 | -104.6256
22.0386 | -104.6349
22.0422 | -104.6453
22.0459 | -104.6569
22.0496 | -104.6673
22.0533 | -104.6743
22.0569 | -104.6813
22.0606 | -104.6879
22.0643 | -104.6947
22.068 | -104.7028
22.0716 | -104.7116
22.0753 | -104.7203
22.079 | -104.7285
22.0826 | -104.7338
22.0863 | -104.7394
22.09 | -104.7445
22.0937 | -104.7467
22.0973 | -104.7501
22.101 | -104.7521
22.1047 | -104.7543
22.1084 | -104.7542
22.112 | -104.752
22.1157 | -104.7507
22.1194 | -104.7508
22.123 | -104.747
22.1267 | -104.7402
22.1304 | -104.7334
22.1341 | -104.7255
22.1377 | -104.7189
22.1414 | -104.7128
22.1451 | -104.7075
22.1488 | -104.7004
22.1524 | -104.6925
22.1561 | -104.688
22.1598 | -104.683
22.1635 | -104.675
22.1671 | -104.6668
22.1708 | -104.6563
22.1745 | -104.6476
22.1781 | -104.6397
22.1818 | -104.6318
22.1855 | -104.6239
22.1892 | -104.6166
22.1928 | -104.6098
22.1965 | -104.6043
22.2002 | -104.5983
22.2039 | -104.5924
22.2075 | -104.5864
22.2112 | -104.5829
22.2149 | -104.5806
22.2185 | -104.577
22.2222 | -104.5756
22.2259 | -104.575
22.2296 | -104.5766
22.2332 | -104.5779
22.2369 | -104.5822
22.2406 | -104.5862
22.2443 | -104.5892
22.2479 | -104.5923
22.2516 | -104.598
22.2553 | -104.6112
22.259 | -104.6172
22.2626 | -104.6251
22.2663 | -104.6337
22.27 | -104.6427
22.2736 | -104.6533
22.2773 | -104.6636
22.281 | -104.6764
22.2847 | -104.689
22.2883 | -104.7013
22.292 | -104.7134
22.2957 | -104.7255
22.2994 | -104.739
22.303 | -104.7506
22.3067 | -104.7638
22.3104 | -104.7788
22.3141 | -104.7933
22.3177 | -104.8068
22.3214 | -104.822
22.3251 | -104.837
22.3287 | -104.8526
22.3324 | -104.8671
22.3361 | -104.8792
22.3398 | -104.891
22.3434 | -104.9064
22.3471 | -104.9206
22.3508 | -104.9349
22.3545 | -104.945
22.3581 | -104.9543
22.3618 | -104.9645
22.3655 | -104.9747
22.3691 | -104.9828
22.3728 | -104.9892
22.3765 | -104.9969
22.3802 | -105.0031
22.3838 | -105.0077
22.3875 | -105.0111
22.3912 | -105.0153
22.3949 | -105.0173
22.3985 | -105.0182
22.4022 | -105.018
22.4059 | -105.0165
22.4095 | -105.0163
22.4132 | -105.0132
22.4169 | -105.0103
22.4206 | -105.0057
22.4242 | -104.9992
22.4279 | -104.9927
22.4316 | -104.985
22.4353 | -104.9753
22.4389 | -104.9641
22.4426 | -104.9531
22.4463 | -104.9414
22.45 | -104.9285
22.4536 | -104.9141
22.4573 | -104.899
22.461 | -104.8845
22.4646 | -104.8691
22.4683 | -104.8517
22.472 | -104.8331
22.4757 | -104.813
22.4793 | -104.7946
22.483 | -104.7779
22.4867 | -104.7622
22.4904 | -104.7475
22.494 | -104.7322
22.4977 | -104.7163
22.5014 | -104.6986
22.5051 | -104.6821
22.5087 | -104.6658
22.5124 | -104.6464
22.5161 | -104.6272
22.5197 | -104.6095
22.5234 | -104.5934
22.5271 | -104.5762
22.5308 | -104.5584
22.5344 | -104.5413
22.5381 | -104.5247
22.5418 | -104.5088
22.5455 | -104.494
22.5491 | -104.4791
22.5528 | -104.4663
22.5565 | -104.4543
22.5601 | -104.4422
22.5638 | -104.4309
22.5675 | -104.4183
22.5712 | -104.4062
22.5748 | -104.3963
22.5785 | -104.3865
22.5822 | -104.3759
22.5859 | -104.366
22.5895 | -104.3588
22.5932 | -104.3532
22.5969 | -104.3487
22.6006 | -104.3427
22.6042 | -104.335
22.6079 | -104.3301
22.6116 | -104.3276
22.6152 | -104.3272
22.6189 | -104.326
22.6226 | -104.3245
22.6263 | -104.3242
22.6299 | -104.3222
22.6336 | -104.3194
22.6373 | -104.321
22.641 | -104.3236
22.6446 | -104.3254
22.6483 | -104.3276
22.652 | -104.3332
22.6556 | -104.3399
22.6593 | -104.3463
22.663 | -104.3515
22.6667 | -104.3581
22.6703 | -104.3654
22.674 | -104.3755
22.6777 | -104.3849
22.6814 | -104.3938
22.685 | -104.4037
22.6887 | -104.4155
22.6924 | -104.4272
22.6961 | -104.4367
22.6997 | -104.4459
22.7034 | -104.4554
22.7071 | -104.4663
22.7107 | -104.4751
22.7144 | -104.4846
22.7181 | -104.4944
22.7218 | -104.5044
22.7254 | -104.5147
22.7291 | -104.5241
22.7328 | -104.5336
22.7365 | -104.5431
22.7401 | -104.5523
22.7438 | -104.5626
22.7475 | -104.5711
22.7511 | -104.5789
22.7548 | -104.5872
22.7585 | -104.598
22.7622 | -104.6076
22.7658 | -104.6138
22.7695 | -104.6192
22.7732 | -104.6259
22.7769 | -104.6309
22.7805 | -104.6349
22.7842 | -104.6389
22.7879 | -104.6441
22.7916 | -104.6484
22.7952 | -104.6522
22.7989 | -104.658
22.8026 | -104.6636
22.8062 | -104.6669
22.8099 | -104.668
22.8136 | -104.6684
22.8173 | -104.6701
22.8209 | -104.6718
22.8246 | -104.6737
22.8283 | -104.6762
22.832 | -104.6789
22.8356 | -104.6813
22.8393 | -104.6828
22.843 | -104.6857
22.8466 | -104.6873
22.8503 | -104.688
22.854 | -104.6887
22.8577 | -104.6875
22.8613 | -104.6856
22.865 | -104.6849
22.8687 | -104.6839
22.8724 | -104.6831
22.876 | -104.6823
22.8797 | -104.6813
22.8834 | -104.682
22.8871 | -104.6812
22.8907 | -104.6798
22.8944 | -104.6779
22.8981 | -104.6765
22.9017 | -104.6761
22.9054 | -104.6758
22.9091 | -104.6759
22.9128 | -104.6721
22.9164 | -104.6692
22.9201 | -104.6672
22.9238 | -104.6656
22.9275 | -104.6661
22.9311 | -104.6642
22.9348 | -104.6614
22.9385 | -104.658
22.9421 | -104.6544
22.9458 | -104.6528
22.9495 | -104.6493
22.9532 | -104.6481
22.9568 | -104.6488
22.9605 | -104.6463
22.9642 | -104.6426
22.9679 | -104.6415
22.9715 | -104.6411
22.9752 | -104.6383
22.9789 | -104.6368
22.9826 | -104.6343
22.9862 | -104.6309
22.9899 | -104.6284
22.9936 | -104.6272
22.9972 | -104.6249
23.0009 | -104.6219
23.0046 | -104.619
23.0083 | -104.617
23.0119 | -104.613
23.0156 | -104.6093
23.0193 | -104.6054
23.023 | -104.6029
23.0266 | -104.6006
23.0303 | -104.5972
23.034 | -104.5926
23.0376 | -104.5872
23.0413 | -104.5821
23.045 | -104.5759
23.0487 | -104.572
23.0523 | -104.5659
23.056 | -104.5612
23.0597 | -104.5577
23.0634 | -104.553
23.067 | -104.5467
23.0707 | -104.5416
23.0744 | -104.5357
23.0781 | -104.53
23.0817 | -104.5251
23.0854 | -104.5198
23.0891 | -104.5156
23.0927 | -104.5123
23.0964 | -104.5072
23.1001 | -104.5032
23.1038 | -104.4983
23.1074 | -104.493
23.1111 | -104.4891
23.1148 | -104.4864
23.1185 | -104.4857
23.1221 | -104.4841
23.1258 | -104.4797
23.1295 | -104.4744
23.1332 | -104.4712
23.1368 | -104.4713
23.1405 | -104.4687
23.1442 | -104.4668
23.1478 | -104.4644
23.1515 | -104.4638
23.1552 | -104.4638
23.1589 | -104.4628
23.1625 | -104.4646
23.1662 | -104.4667
23.1699 | -104.4646
23.1736 | -104.4618
23.1772 | -104.4626
23.1809 | -104.4652
23.1846 | -104.469
23.1882 | -104.4703
23.1919 | -104.4677
23.1956 | -104.4651
23.1993 | -104.4643
23.2029 | -104.4624
23.2066 | -104.462
23.2103 | -104.461
23.214 | -104.4601
23.2176 | -104.4597
23.2213 | -104.4596
23.225 | -104.4592
23.2286 | -104.4581
23.2323 | -104.4571
23.236 | -104.455
23.2397 | -104.4534
23.2433 | -104.4515
23.247 | -104.4478
23.2507 | -104.4429
23.2544 | -104.4381
23.258 | -104.434
23.2617 | -104.4309
23.2654 | -104.4265
23.2691 | -104.4222
23.2727 | -104.4175
23.2764 | -104.4141
23.2801 | -104.4087
23.2837 | -104.4
23.2874 | -104.3921
23.2911 | -104.3852
23.2948 | -104.3764
23.2984 | -104.3685
23.3021 | -104.3613
23.3058 | -104.3542
23.3095 | -104.3491
23.3131 | -104.3403
23.3168 | -104.334
23.3205 | -104.3257
23.3242 | -104.3159
23.3278 | -104.3062
23.3315 | -104.2981
23.3352 | -104.2896
23.3388 | -104.279
23.3425 | -104.2699
23.3462 | -104.2615
23.3499 | -104.2555
23.3535 | -104.2473
23.3572 | -104.2383
23.3609 | -104.2301
23.3646 | -104.2242
23.3682 | -104.2168
23.3719 | -104.2082
23.3756 | -104.2
23.3792 | -104.1883
23.3829 | -104.1791
23.3866 | -104.1687
23.3903 | -104.1592
23.3939 | -104.1515
23.3976 | -104.1437
23.4013 | -104.1358
23.405 | -104.1286
23.4086 | -104.121
23.4123 | -104.1102
23.416 | -104.0995
23.4197 | -104.0919
23.4233 | -104.0843
23.427 | -104.076
23.4307 | -104.067
23.4343 | -104.0596
23.438 | -104.0514
23.4417 | -104.0415
23.4454 | -104.0321
23.449 | -104.0241
23.4527 | -104.0171
23.4564 | -104.0123
23.4601 | -104.0069
23.4637 | -103.9999
23.4674 | -103.9914
23.4711 | -103.9853
23.4747 | -103.9784
23.4784 | -103.9748
23.4821 | -103.9696
23.4858 | -103.9646
23.4894 | -103.96
23.4931 | -103.9552
23.4968 | -103.9508
23.5005 | -103.949
23.5041 | -103.9484
23.5078 | -103.946
23.5115 | -103.9425
23.5152 | -103.9395
23.5188 | -103.9387
23.5225 | -103.9399
23.5262 | -103.9408
23.5298 | -103.941
23.5335 | -103.9417
23.5372 | -103.9416
23.5409 | -103.9426
23.5445 | -103.9461
23.5482 | -103.9496
23.5519 | -103.9534
23.5556 | -103.9565
23.5592 | -103.9598
23.5629 | -103.9638
23.5666 | -103.9696
23.5702 | -103.9783
23.5739 | -103.984
23.5776 | -103.991
23.5813 | -104.0007
23.5849 | -104.0088
23.5886 | -104.0181
23.5923 | -104.0278
23.596 | -104.0391
23.5996 | -104.0501
23.6033 | -104.0622
23.607 | -104.0746
23.6107 | -104.0876
23.6143 | -104.1044
23.618 | -104.122
23.6217 | -104.14
23.6253 | -104.1582
23.629 | -104.1756
23.6327 | -104.1919
23.6364 | -104.2089
23.64 | -104.2269
23.6437 | -104.2471
23.6474 | -104.2699
23.6511 | -104.2917
23.6547 | -104.3155
23.6584 | -104.3403
23.6621 | -104.365
23.6657 | -104.3889
23.6694 | -104.4133
23.6731 | -104.4381
23.6768 | -104.4621
23.6804 | -104.4854
23.6841 | -104.5056
23.6878 | -104.5277
23.6915 | -104.5494
23.6951 | -104.5706
23.6988 | -104.5917
23.7025 | -104.6132
23.7062 | -104.6346
23.7098 | -104.6547
23.7135 | -104.6753
23.7172 | -104.6931
23.7208 | -104.7102
23.7245 | -104.7233
23.7282 | -104.7343
23.7319 | -104.7466
23.7355 | -104.7593
23.7392 | -104.7708
23.7429 | -104.7796
23.7466 | -104.7877
23.7502 | -104.7956
23.7539 | -104.8007
23.7576 | -104.8058
23.7612 | -104.8113
23.7649 | -104.8154
23.7686 | -104.818
23.7723 | -104.8205
23.7759 | -104.822
23.7796 | -104.8229
23.7833 | -104.8201
23.787 | -104.817
23.7906 | -104.8116
23.7943 | -104.8072
23.798 | -104.8039
23.8017 | -104.8003
23.8053 | -104.7968
23.809 | -104.7932
23.8127 | -104.786
23.8163 | -104.7793
23.82 | -104.773
23.8237 | -104.767
23.8274 | -104.7614
23.831 | -104.7574
23.8347 | -104.7527
23.8384 | -104.7465
23.8421 | -104.7426
23.8457 | -104.7458
23.8494 | -104.744
23.8531 | -104.7431
23.8567 | -104.7425
23.8604 | -104.7423
23.8641 | -104.7428
23.8678 | -104.7428
23.8714 | -104.7437
23.8751 | -104.7458
23.8788 | -104.7469
23.8825 | -104.7481
23.8861 | -104.7497
23.8898 | -104.7515
23.8935 | -104.7552
23.8972 | -104.7599
23.9008 | -104.7665
23.9045 | -104.7726
23.9082 | -104.7777
23.9118 | -104.7855
23.9155 | -104.7934
23.9192 | -104.7998
23.9229 | -104.8084
23.9265 | -104.816
23.9302 | -104.8233
23.9339 | -104.8309
23.9376 | -104.8393
23.9412 | -104.8467
23.9449 | -104.8549
23.9486 | -104.8639
23.9522 | -104.8729
23.9559 | -104.8804
23.9596 | -104.8877
23.9633 | -104.8957
23.9669 | -104.9052
23.9706 | -104.916
23.9743 | -104.9252
23.978 | -104.9329
23.9816 | -104.9407
23.9853 | -104.9464
23.989 | -104.9518
23.9927 | -104.9549
23.9963 | -104.9584
24 | -104.9603
24.0037 | -104.9635
24.0073 | -104.9672
24.011 | -104.9713
24.0147 | -104.9772
24.0184 | -104.9831
24.022 | -104.9883
24.0257 | -104.9913
24.0294 | -104.9941
24.0331 | -104.9959
24.0367 | -104.9969
24.0404 | -105.0004
24.0441 | -105.0038
24.0478 | -105.0074
24.0514 | -105.0098
24.0551 | -105.0113
24.0588 | -105.0113
24.0624 | -105.013
24.0661 | -105.0169
24.0698 | -105.0188
24.0735 | -105.0185
24.0771 | -105.0166
24.0808 | -105.0184
24.0845 | -105.0235
24.0882 | -105.0263
24.0918 | -105.0273
24.0955 | -105.0267
24.0992 | -105.0266
24.1028 | -105.0265
24.1065 | -105.0249
24.1102 | -105.0236
24.1139 | -105.0245
24.1175 | -105.0253
24.1212 | -105.0266
24.1249 | -105.0281
24.1286 | -105.0316
24.1322 | -105.0343
24.1359 | -105.0354
24.1396 | -105.0377
24.1433 | -105.0401
24.1469 | -105.0423
24.1506 | -105.0454
24.1543 | -105.0492
24.1579 | -105.0524
24.1616 | -105.0568
24.1653 | -105.062
24.169 | -105.0662
24.1726 | -105.0723
24.1763 | -105.0798
24.18 | -105.0858
24.1837 | -105.0917
24.1873 | -105.0963
24.191 | -105.1021
24.1947 | -105.1078
24.1983 | -105.1142
24.202 | -105.1222
24.2057 | -105.1318
24.2094 | -105.1421
24.213 | -105.1507
24.2167 | -105.1596
24.2204 | -105.1688
24.2241 | -105.1771
24.2277 | -105.1863
24.2314 | -105.1956
24.2351 | -105.2041
24.2388 | -105.2141
24.2424 | -105.2217
24.2461 | -105.2281
24.2498 | -105.2343
24.2534 | -105.2399
24.2571 | -105.2463
24.2608 | -105.2551
24.2645 | -105.2623
24.2681 | -105.2647
24.2718 | -105.2657
24.2755 | -105.2681
24.2792 | -105.2708
24.2828 | -105.2689
24.2865 | -105.2658
24.2902 | -105.2615
24.2938 | -105.2543
24.2975 | -105.2481
24.3012 | -105.2398
24.3049 | -105.2304
24.3085 | -105.2201
24.3122 | -105.2049
24.3159 | -105.1874
24.3196 | -105.1728
24.3232 | -105.1557
24.3269 | -105.1407
24.3306 | -105.1255
24.3343 | -105.1097
24.3379 | -105.0916
24.3416 | -105.0733
24.3453 | -105.0511
24.3489 | -105.0253
24.3526 | -105.0001
24.3563 | -104.9724
24.36 | -104.9431
24.3636 | -104.9143
24.3673 | -104.8848
24.371 | -104.8538
24.3747 | -104.8146
24.3783 | -104.7838
24.382 | -104.7534
24.3857 | -104.7221
24.3893 | -104.6893
24.393 | -104.6578
24.3967 | -104.6285
24.4004 | -104.6001
24.404 | -104.5731
24.4077 | -104.5462
24.4114 | -104.5182
24.4151 | -104.4921
24.4187 | -104.4684
24.4224 | -104.4468
24.4261 | -104.4247
24.4298 | -104.4034
24.4334 | -104.3858
24.4371 | -104.3696
24.4408 | -104.3556
24.4444 | -104.3433
24.4481 | -104.3335
24.4518 | -104.3265
24.4555 | -104.321
24.4591 | -104.3169
24.4628 | -104.3149
24.4665 | -104.3156
24.4702 | -104.3173
24.4738 | -104.321
24.4775 | -104.3265
24.4812 | -104.3357
24.4848 | -104.3475
24.4885 | -104.3603
24.4922 | -104.3745
24.4959 | -104.3913
24.4995 | -104.41
24.5032 | -104.4302
24.5069 | -104.451
24.5106 | -104.4725
24.5142 | -104.496
24.5179 | -104.5211
24.5216 | -104.5471
24.5253 | -104.5743
24.5289 | -104.6015
24.5326 | -104.6309
24.5363 | -104.6598
24.5399 | -104.6898
24.5436 | -104.72
24.5473 | -104.7479
24.551 | -104.7745
24.5546 | -104.8018
24.5583 | -104.8279
24.562 | -104.8524
24.5657 | -104.876
24.5693 | -104.8956
24.573 | -104.9137
24.5767 | -104.9278
24.5803 | -104.9306
24.584 | -104.9293
24.5877 | -104.925
24.5914 | -104.9133
24.595 | -104.8932
24.5987 | -104.8637
24.6024 | -104.8266
24.6061 | -104.7908
24.6097 | -104.7562
24.6134 | -104.7216
24.6171 | -104.6867
24.6208 | -104.6434
24.6244 | -104.5963
24.6281 | -104.5488
24.6318 | -104.4979
24.6354 | -104.4434
24.6391 | -104.384
24.6428 | -104.3206
24.6465 | -104.2561
24.6501 | -104.1901
24.6538 | -104.1206
24.6575 | -104.0515
24.6612 | -103.9799
24.6648 | -103.9031
24.6685 | -103.8245
24.6722 | -103.7439
24.6758 | -103.6618
24.6795 | -103.5793
24.6832 | -103.496
24.6869 | -103.4132
24.6905 | -103.3323
24.6942 | -103.2531
24.6979 | -103.1749
24.7016 | -103.0979
24.7052 | -103.0237
24.7089 | -102.9517
24.7126 | -102.8828
24.7163 | -102.8183
24.7199 | -102.7568
24.7236 | -102.698
24.7273 | -102.6428
24.7309 | -102.5908
24.7346 | -102.5428
24.7383 | -102.4995
24.742 | -102.4594
24.7456 | -102.4237
24.7493 | -102.3941
24.753 | -102.37
24.7567 | -102.3502
24.7603 | -102.3347
24.764 | -102.3253
24.7677 | -102.3206
24.7714 | -102.3214
24.775 | -102.3275
24.7787 | -102.3381
24.7824 | -102.3535
24.786 | -102.3748
24.7897 | -102.4017
24.7934 | -102.4341
24.7971 | -102.4713
24.8007 | -102.5128
24.8044 | -102.5588
24.8081 | -102.6094
24.8118 | -102.6636
24.8154 | -102.7218
24.8191 | -102.7849
24.8228 | -102.8508
24.8264 | -102.9195
24.8301 | -102.9914
24.8338 | -103.0653
24.8375 | -103.142
24.8411 | -103.2218
24.8448 | -103.3026
24.8485 | -103.3845
24.8522 | -103.4671
24.8558 | -103.5509
24.8595 | -103.6348
24.8632 | -103.7189
24.8668 | -103.8029
24.8705 | -103.8866
24.8742 | -103.9681
24.8779 | -104.0491
24.8815 | -104.1287
24.8852 | -104.2067
24.8889 | -104.2817
24.8926 | -104.3543
24.8962 | -104.4255
24.8999 | -104.4936
24.9036 | -104.5257
24.9073 | -104.587
24.9109 | -104.6475
24.9146 | -104.7055
24.9183 | -104.7614
24.9219 | -104.8168
24.9256 | -104.8677
24.9293 | -104.9174
24.933 | -104.9693
24.9366 | -105.0178
24.9403 | -105.0636
24.944 | -105.11
24.9477 | -105.1546
24.9513 | -105.201
24.955 | -105.2444
24.9587 | -105.2728
24.9624 | -105.2931
24.966 | -105.3088
24.9697 | -105.3198
24.9734 | -105.3321
24.977 | -105.3423
24.9807 | -105.3504
24.9844 | -105.3538
24.9881 | -105.3556
24.9917 | -105.3589
24.9954 | -105.3597
24.9991 | -105.3592
25.0028 | -105.3565
25.0064 | -105.3521
25.0101 | -105.3487
25.0138 | -105.346
25.0174 | -105.3448
25.0211 | -105.3438
25.0248 | -105.3418
25.0285 | -105.3383
25.0321 | -105.3358
25.0358 | -105.3342
25.0395 | -105.331
25.0432 | -105.3255
25.0468 | -105.3209
25.0505 | -105.3151
25.0542 | -105.3082
25.0579 | -105.3002
25.0615 | -105.2923
25.0652 | -105.2834
25.0689 | -105.2737
25.0725 | -105.2641
25.0762 | -105.2539
25.0799 | -105.2435
25.0836 | -105.2338
25.0872 | -105.2244
25.0909 | -105.2149
25.0946 | -105.2055
25.0983 | -105.1954
25.1019 | -105.1854
25.1056 | -105.173
25.1093 | -105.1605
25.1129 | -105.1499
25.1166 | -105.1373
25.1203 | -105.1256
25.124 | -105.1132
25.1276 | -105.1017
25.1313 | -105.0901
25.135 | -105.0792
25.1387 | -105.0679
25.1423 | -105.0577
25.146 | -105.0463
25.1497 | -105.0371
25.1534 | -105.0275
25.157 | -105.0171
25.1607 | -105.0056
25.1644 | -104.9956
25.168 | -104.9872
25.1717 | -104.9805
25.1754 | -104.9738
25.1791 | -104.9689
25.1827 | -104.9629
25.1864 | -104.9571
25.1901 | -104.9512
25.1938 | -104.9442
25.1974 | -104.9366
25.2011 | -104.9309
25.2048 | -104.9274
25.2084 | -104.9249
25.2121 | -104.9226
25.2158 | -104.9192
25.2195 | -104.9158
25.2231 | -104.9134
25.2268 | -104.9125
25.2305 | -104.9125
25.2342 | -104.9108
25.2378 | -104.909
25.2415 | -104.9082
25.2452 | -104.9076
25.2489 | -104.9068
25.2525 | -104.9081
25.2562 | -104.9097
25.2599 | -104.9101
25.2635 | -104.9101
25.2672 | -104.9095
25.2709 | -104.9093
25.2746 | -104.9081
25.2782 | -104.907
25.2819 | -104.9075
25.2856 | -104.9082
25.2893 | -104.9116
25.2929 | -104.9144
25.2966 | -104.9162
25.3003 | -104.9182
25.3039 | -104.9207
25.3076 | -104.9219
25.3113 | -104.9235
25.315 | -104.9264
25.3186 | -104.9286
25.3223 | -104.9317
25.326 | -104.9335
25.3297 | -104.9366
25.3333 | -104.9378
25.337 | -104.9386
25.3407 | -104.9408
25.3444 | -104.9431
25.348 | -104.9437
25.3517 | -104.9441
25.3554 | -104.9457
25.359 | -104.9486
25.3627 | -104.9513
25.3664 | -104.9529
25.3701 | -104.9541
25.3737 | -104.954
25.3774 | -104.954
25.3811 | -104.9541
25.3848 | -104.9543
25.3884 | -104.9537
25.3921 | -104.953
25.3958 | -104.9541
25.3994 | -104.9553
25.4031 | -104.9543
25.4068 | -104.9539
25.4105 | -104.9533
25.4141 | -104.952
25.4178 | -104.9513
25.4215 | -104.9518
25.4252 | -104.9501
25.4288 | -104.9466
25.4325 | -104.9438
25.4362 | -104.943
25.4399 | -104.9417
25.4435 | -104.941
25.4472 | -104.9389
25.4509 | -104.9378
25.4545 | -104.9362
25.4582 | -104.9343
25.4619 | -104.9312
25.4656 | -104.9283
25.4692 | -104.9259
25.4729 | -104.9242
25.4766 | -104.9215
25.4803 | -104.9206
25.4839 | -104.9201
25.4876 | -104.9177
25.4913 | -104.9175
25.4949 | -104.9176
25.4986 | -104.9176
25.5023 | -104.9151
25.506 | -104.9129
25.5096 | -104.9115
25.5133 | -104.9109
25.517 | -104.9114
25.5207 | -104.9127
25.5243 | -104.9136
25.528 | -104.9126
25.5317 | -104.9099
25.5354 | -104.9089
25.539 | -104.9074
25.5427 | -104.9061
25.5464 | -104.9059
25.55 | -104.9057
25.5537 | -104.9056
25.5574 | -104.906
25.5611 | -104.9048
25.5647 | -104.9023
25.5684 | -104.9002
25.5721 | -104.8998
25.5758 | -104.8973
25.5794 | -104.8949
25.5831 | -104.8932
25.5868 | -104.8913
25.5905 | -104.8888
25.5941 | -104.8892
25.5978 | -104.8895
25.6015 | -104.889
25.6051 | -104.888
25.6088 | -104.887
25.6125 | -104.8834
25.6162 | -104.8794
25.6198 | -104.8788
25.6235 | -104.879
25.6272 | -104.8775
25.6309 | -104.8731
25.6345 | -104.8708
25.6382 | -104.8687
25.6419 | -104.8648
25.6455 | -104.861
25.6492 | -104.8552
25.6529 | -104.8495
25.6566 | -104.8445
25.6602 | -104.8402
25.6639 | -104.8365
25.6676 | -104.8326
25.6713 | -104.8284
25.6749 | -104.8263
25.6786 | -104.8243
25.6823 | -104.8204
25.6859 | -104.8178
25.6896 | -104.814
25.6933 | -104.812
25.697 | -104.8096
25.7006 | -104.8078
25.7043 | -104.8067
25.708 | -104.8072
25.7117 | -104.8073
25.7153 | -104.8046
25.719 | -104.8033
25.7227 | -104.8035
25.7264 | -104.8033
25.73 | -104.8043
25.7337 | -104.8051
25.7374 | -104.8066
25.741 | -104.8071
25.7447 | -104.8058
25.7484 | -104.8076
25.7521 | -104.8117
25.7557 | -104.8154
25.7594 | -104.8195
25.7631 | -104.8233
25.7668 | -104.8281
25.7704 | -104.8344
25.7741 | -104.8411
25.7778 | -104.846
25.7815 | -104.8505
25.7851 | -104.8589
25.7888 | -104.8672
25.7925 | -104.8761
25.7961 | -104.8829
25.7998 | -104.8904
25.8035 | -104.8987
25.8072 | -104.9077
25.8108 | -104.9162
25.8145 | -104.9249
25.8182 | -104.931
25.8219 | -104.9392
25.8255 | -104.9499
25.8292 | -104.9613
25.8329 | -104.9735
25.8365 | -104.9835
25.8402 | -104.9946
25.8439 | -105.0043
25.8476 | -105.0132
25.8512 | -105.0222
25.8549 | -105.0297
25.8586 | -105.0367
25.8623 | -105.0438
25.8659 | -105.0535
25.8696 | -105.0634
25.8733 | -105.0697
25.877 | -105.0758
25.8806 | -105.0834
25.8843 | -105.0912
25.888 | -105.0993
25.8916 | -105.1065
25.8953 | -105.1123
25.899 | -105.1171
25.9027 | -105.1222
25.9063 | -105.1273
25.91 | -105.132
25.9137 | -105.1366
25.9174 | -105.1402
25.921 | -105.1422
25.9247 | -105.1458
25.9284 | -105.1481
25.932 | -105.1491
25.9357 | -105.1487
25.9394 | -105.151
25.9431 | -105.1529
25.9467 | -105.154
25.9504 | -105.155
25.9541 | -105.1565
25.9578 | -105.1549
25.9614 | -105.1528
25.9651 | -105.1538
25.9688 | -105.1536
25.9725 | -105.1535
25.9761 | -105.1555
25.9798 | -105.1586
25.9835 | -105.1613
25.9871 | -105.1627
25.9908 | -105.1634
25.9945 | -105.1668
25.9982 | -105.172
26.0018 | -105.1755
26.0055 | -105.1793
26.0092 | -105.1842
26.0129 | -105.1893
26.0165 | -105.196
26.0202 | -105.2013
26.0239 | -105.205
26.0275 | -105.2103
26.0312 | -105.2174
26.0349 | -105.2234
26.0386 | -105.2287
26.0422 | -105.2342
26.0459 | -105.2405
26.0496 | -105.2472
26.0533 | -105.2532
26.0569 | -105.2589
26.0606 | -105.2653
26.0643 | -105.2735
26.068 | -105.2823
26.0716 | -105.2901
26.0753 | -105.2964
26.079 | -105.3036
26.0826 | -105.3107
26.0863 | -105.3172
26.09 | -105.3232
26.0937 | -105.3285
26.0973 | -105.3321
26.101 | -105.337
26.1047 | -105.342
26.1084 | -105.3488
26.112 | -105.3574
26.1157 | -105.3626
26.1194 | -105.3645
26.123 | -105.3656
26.1267 | -105.3683
26.1304 | -105.3724
26.1341 | -105.3754
26.1377 | -105.3775
26.1414 | -105.3797
26.1451 | -105.382
26.1488 | -105.3846
26.1524 | -105.3854
26.1561 | -105.3882
26.1598 | -105.3897
26.1635 | -105.3915
26.1671 | -105.3933
26.1708 | -105.3956
26.1745 | -105.3963
26.1781 | -105.3989
26.1818 | -105.401
26.1855 | -105.4069
26.1892 | -105.4115
26.1928 | -105.4126
26.1965 | -105.4124
26.2002 | -105.4129
26.2039 | -105.4146
26.2075 | -105.4152
26.2112 | -105.414
26.2149 | -105.4136
26.2185 | -105.4155
26.2222 | -105.4172
26.2259 | -105.4175
26.2296 | -105.4175
26.2332 | -105.417
26.2369 | -105.4163
26.2406 | -105.4187
26.2443 | -105.421
26.2479 | -105.4224
26.2516 | -105.4218
26.2553 | -105.422
26.259 | -105.4247
26.2626 | -105.4266
26.2663 | -105.4284
26.27 | -105.4297
26.2736 | -105.4308
26.2773 | -105.4318
26.281 | -105.4327
26.2847 | -105.4348
26.2883 | -105.4383
26.292 | -105.4394
26.2957 | -105.4415
26.2994 | -105.4437
26.303 | -105.4455
26.3067 | -105.4476
26.3104 | -105.4485
26.3141 | -105.4488
26.3177 | -105.4481
26.3214 | -105.4456
26.3251 | -105.4456
26.3287 | -105.4468
26.3324 | -105.4466
26.3361 | -105.4476
26.3398 | -105.4481
26.3434 | -105.4484
26.3471 | -105.4464
26.3508 | -105.445
26.3545 | -105.4453
26.3581 | -105.4444
26.3618 | -105.4436
26.3655 | -105.4422
26.3691 | -105.441
26.3728 | -105.4416
26.3765 | -105.4403
26.3802 | -105.4384
26.3838 | -105.436
26.3875 | -105.4345
26.3912 | -105.4339
26.3949 | -105.4345
26.3985 | -105.4348
26.4022 | -105.4327
26.4059 | -105.4316
26.4095 | -105.431
26.4132 | -105.4302
26.4169 | -105.4309
26.4206 | -105.4331
26.4242 | -105.4342
26.4279 | -105.4356
26.4316 | -105.4377
26.4353 | -105.4399
26.4389 | -105.4411
26.4426 | -105.4422
26.4463 | -105.4445
26.45 | -105.4476
26.4536 | -105.4517
26.4573 | -105.4552
26.461 | -105.4577
26.4646 | -105.4613
26.4683 | -105.4644
26.472 | -105.4669
26.4757 | -105.4695
26.4793 | -105.472
26.483 | -105.4757
26.4867 | -105.4777
26.4904 | -105.4803
26.494 | -105.4823
26.4977 | -105.4848
26.5014 | -105.4869
26.5051 | -105.4886
26.5087 | -105.4897
26.5124 | -105.4917
26.5161 | -105.4935
26.5197 | -105.4967
26.5234 | -105.4988
26.5271 | -105.4987
26.5308 | -105.4989
26.5344 | -105.5002
26.5381 | -105.5005
26.5418 | -105.4994
26.5455 | -105.4982
26.5491 | -105.4979
26.5528 | -105.497
26.5565 | -105.4959
26.5601 | -105.4953
26.5638 | -105.4934
26.5675 | -105.4926
26.5712 | -105.4941
26.5748 | -105.4936
26.5785 | -105.4924
26.5822 | -105.4905
26.5859 | -105.4894
26.5895 | -105.4878
26.5932 | -105.4859
26.5969 | -105.4843
26.6006 | -105.4839
26.6042 | -105.4829
26.6079 | -105.4812
26.6116 | -105.4792
26.6152 | -105.4764
26.6189 | -105.4742
26.6226 | -105.4717
26.6263 | -105.4698
26.6299 | -105.4669
26.6336 | -105.4627
26.6373 | -105.4597
26.641 | -105.4573
26.6446 | -105.4554
26.6483 | -105.453
26.652 | -105.4518
26.6556 | -105.449
26.6593 | -105.4462
26.663 | -105.4434
26.6667 | -105.4416
26.6703 | -105.4401
26.674 | -105.4378
26.6777 | -105.4362
26.6814 | -105.4328
26.685 | -105.4307
26.6887 | -105.4295
26.6924 | -105.4278
26.6961 | -105.4271
26.6997 | -105.4268
26.7034 | -105.4261
26.7071 | -105.4244
26.7107 | -105.4209
26.7144 | -105.4197
26.7181 | -105.418
26.7218 | -105.4168
26.7254 | -105.4167
26.7291 | -105.4147
26.7328 | -105.4111
26.7365 | -105.4087
26.7401 | -105.4066
26.7438 | -105.4058
26.7475 | -105.4042
26.7511 | -105.4048
26.7548 | -105.4052
26.7585 | -105.4058
26.7622 | -105.4038
26.7658 | -105.4029
26.7695 | -105.4037
26.7732 | -105.4034
26.7769 | -105.4047
26.7805 | -105.4074
26.7842 | -105.4093
26.7879 | -105.41
26.7916 | -105.4117
26.7952 | -105.4126
26.7989 | -105.4144
26.8026 | -105.4168
26.8062 | -105.4175
26.8099 | -105.4164
26.8136 | -105.4163
26.8173 | -105.4178
26.8209 | -105.4199
26.8246 | -105.4205
26.8283 | -105.4199
26.832 | -105.4198
26.8356 | -105.4206
26.8393 | -105.4222
26.843 | -105.4244
26.8466 | -105.4257
26.8503 | -105.424
26.854 | -105.4214
26.8577 | -105.42
26.8613 | -105.4199
26.865 | -105.4198
26.8687 | -105.4195
26.8724 | -105.4176
26.876 | -105.4163
26.8797 | -105.4155
26.8834 | -105.4131
26.8871 | -105.4109
26.8907 | -105.4086
26.8944 | -105.4066
26.8981 | -105.4039
26.9017 | -105.4013
26.9054 | -105.399
26.9091 | -105.3965
26.9128 | -105.3921
26.9164 | -105.3896
26.9201 | -105.3881
26.9238 | -105.385
26.9275 | -105.3826
26.9311 | -105.3792
26.9348 | -105.375
26.9385 | -105.3723
26.9421 | -105.3704
26.9458 | -105.3663
26.9495 | -105.3609
26.9532 | -105.3562
26.9568 | -105.3523
26.9605 | -105.3493
26.9642 | -105.3466
26.9679 | -105.3436
26.9715 | -105.3404
26.9752 | -105.3371
26.9789 | -105.3318
26.9826 | -105.3276
26.9862 | -105.324
26.9899 | -105.321
26.9936 | -105.3179
26.9972 | -105.3145
27.0009 | -105.3111
27.0046 | -105.3067
27.0083 | -105.3033
27.0119 | -105.3013
27.0156 | -105.2985
27.0193 | -105.2955
27.023 | -105.2924
27.0266 | -105.2887
27.0303 | -105.2845
27.034 | -105.2792
27.0376 | -105.2768
27.0413 | -105.2749
27.045 | -105.2726
27.0487 | -105.2686
27.0523 | -105.2647
27.056 | -105.2615
27.0597 | -105.2582
27.0634 | -105.2554
27.067 | -105.2526
27.0707 | -105.249
27.0744 | -105.2465
27.0781 | -105.2438
27.0817 | -105.2379
27.0854 | -105.2325
27.0891 | -105.2284
27.0927 | -105.2258
27.0964 | -105.222
27.1001 | -105.218
27.1038 | -105.2141
27.1074 | -105.2087
27.1111 | -105.2034
27.1148 | -105.1996
27.1185 | -105.195
27.1221 | -105.1902
27.1258 | -105.1839
27.1295 | -105.1781
27.1332 | -105.172
27.1368 | -105.1662
27.1405 | -105.1593
27.1442 | -105.1539
27.1478 | -105.1483
27.1515 | -105.1414
27.1552 | -105.1347
27.1589 | -105.129
27.1625 | -105.1232
27.1662 | -105.1171
27.1699 | -105.1113
27.1736 | -105.1051
27.1772 | -105.1009
27.1809 | -105.097
27.1846 | -105.0937
27.1882 | -105.09
27.1919 | -105.0873
27.1956 | -105.0832
27.1993 | -105.0787
27.2029 | -105.0741
27.2066 | -105.0698
27.2103 | -105.0667
27.214 | -105.0635
27.2176 | -105.0594
27.2213 | -105.0551
27.225 | -105.0522
27.2286 | -105.049
27.2323 | -105.0475
27.236 | -105.0462
27.2397 | -105.0433
27.2433 | -105.0395
27.247 | -105.0361
27.2507 | -105.0334
27.2544 | -105.0318
27.258 | -105.0305
27.2617 | -105.0293
27.2654 | -105.0287
27.2691 | -105.0276
27.2727 | -105.0269
27.2764 | -105.026
27.2801 | -105.0247
27.2837 | -105.0236
27.2874 | -105.0243
27.2911 | -105.025
27.2948 | -105.0258
27.2984 | -105.026
27.3021 | -105.0272
27.3058 | -105.0281
27.3095 | -105.0271
27.3131 | -105.0266
27.3168 | -105.027
27.3205 | -105.0274
27.3242 | -105.0286
27.3278 | -105.0309
27.3315 | -105.0336
27.3352 | -105.0371
27.3388 | -105.0402
27.3425 | -105.0441
27.3462 | -105.0486
27.3499 | -105.0526
27.3535 | -105.0544
27.3572 | -105.0568
27.3609 | -105.0607
27.3646 | -105.0644
27.3682 | -105.0691
27.3719 | -105.0741
27.3756 | -105.0806
27.3792 | -105.086
27.3829 | -105.0899
27.3866 | -105.0958
27.3903 | -105.1024
27.3939 | -105.1081
27.3976 | -105.1147
27.4013 | -105.1209
27.405 | -105.1267
27.4086 | -105.1303
27.4123 | -105.1352
27.416 | -105.1416
27.4197 | -105.1469
27.4233 | -105.1507
27.427 | -105.1547
27.4307 | -105.16
27.4343 | -105.1658
27.438 | -105.1702
27.4417 | -105.1755
27.4454 | -105.1793
27.449 | -105.1829
27.4527 | -105.1869
27.4564 | -105.1909
27.4601 | -105.1947
27.4637 | -105.1986
27.4674 | -105.2038
27.4711 | -105.2084
27.4747 | -105.211
27.4784 | -105.2112
27.4821 | -105.2106
27.4858 | -105.2105
27.4894 | -105.2108
27.4931 | -105.2115
27.4968 | -105.214
27.5005 | -105.2161
27.5041 | -105.2175
27.5078 | -105.2189
27.5115 | -105.2201
27.5152 | -105.2195
27.5188 | -105.2196
27.5225 | -105.2187
27.5262 | -105.2191
27.5298 | -105.219
27.5335 | -105.2175
27.5372 | -105.217
27.5409 | -105.2168
27.5445 | -105.2142
27.5482 | -105.2093
27.5519 | -105.2069
27.5556 | -105.2064
27.5592 | -105.2053
27.5629 | -105.204
27.5666 | -105.2019
27.5702 | -105.2001
27.5739 | -105.1981
27.5776 | -105.1972
27.5813 | -105.197
27.5849 | -105.1975
27.5886 | -105.1977
27.5923 | -105.1956
27.596 | -105.1933
27.5996 | -105.1921
27.6033 | -105.1913
27.607 | -105.1902
27.6107 | -105.189
27.6143 | -105.1868
27.618 | -105.1847
27.6217 | -105.1826
27.6253 | -105.1808
27.629 | -105.1809
27.6327 | -105.1819
27.6364 | -105.1809
27.64 | -105.1781
27.6437 | -105.1765
27.6474 | -105.1755
27.6511 | -105.1736
27.6547 | -105.1726
27.6584 | -105.1716
27.6621 | -105.1717
27.6657 | -105.1729
27.6694 | -105.1739
27.6731 | -105.1742
27.6768 | -105.1746
27.6804 | -105.1745
27.6841 | -105.175
27.6878 | -105.1758
27.6915 | -105.1758
27.6951 | -105.1771
27.6988 | -105.1795
27.7025 | -105.1826
27.7062 | -105.1857
27.7098 | -105.1871
27.7135 | -105.1886
27.7172 | -105.1905
27.7208 | -105.1933
27.7245 | -105.196
27.7282 | -105.1992
27.7319 | -105.2022
27.7355 | -105.2051
27.7392 | -105.2084
27.7429 | -105.2135
27.7466 | -105.2175
27.7502 | -105.2208
27.7539 | -105.2249
27.7576 | -105.2292
27.7612 | -105.2333
27.7649 | -105.2365
27.7686 | -105.2424
27.7723 | -105.2494
27.7759 | -105.2569
27.7796 | -105.2625
27.7833 | -105.2691
27.787 | -105.2755
27.7906 | -105.2796
27.7943 | -105.2863
27.798 | -105.2953
27.8017 | -105.3009
27.8053 | -105.3006
27.809 | -105.3037
27.8127 | -105.3104
27.8163 | -105.3165
27.82 | -105.322
27.8237 | -105.3281
27.8274 | -105.3338
27.831 | -105.3389
27.8347 | -105.3453
27.8384 | -105.3495
27.8421 | -105.3529
27.8457 | -105.3582
27.8494 | -105.3644
27.8531 | -105.3695
27.8567 | -105.3749
27.8604 | -105.3795
27.8641 | -105.3845
27.8678 | -105.3897
27.8714 | -105.3958
27.8751 | -105.4017
27.8788 | -105.4053
27.8825 | -105.4065
27.8861 | -105.4069
27.8898 | -105.4094
27.8935 | -105.4093
27.8972 | -105.4079
27.9008 | -105.4074
27.9045 | -105.4068
27.9082 | -105.4054
27.9118 | -105.4044
27.9155 | -105.404
27.9192 | -105.4032
27.9229 | -105.4017
27.9265 | -105.3971
27.9302 | -105.3914
27.9339 | -105.3884
27.9376 | -105.3866
27.9412 | -105.3829
27.9449 | -105.3764
27.9486 | -105.3703
27.9522 | -105.3629
27.9559 | -105.354
27.9596 | -105.3448
27.9633 | -105.3371
27.9669 | -105.3283
27.9706 | -105.317
27.9743 | -105.305
27.978 | -105.2947
27.9816 | -105.2859
27.9853 | -105.2778
27.989 | -105.27
27.9927 | -105.26
27.9963 | -105.2482
28 | -105.2366
28.0037 | -105.2245