## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3d | 29 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 21
Hood | 02-91C/P | 19
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1001.2 | hPa |  | PRC PRO 01
Humidity | 31.6 | %/Hr
Internal temperature of calibration bench | 23.44 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1008 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 23.9 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.84 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.89 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3233 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Analog acquisition system | Analog Hood
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 01/03/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 16/03/2018
Technical Manager | Anthony HUE | Approval |  | 20/03/2018
CEO | Robert GLEMOT | Approval |  | 20/03/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3926 | 57.4407
0.012209 | -35.9502 | 38.4
0.018313 | -34.9089 | 28.4498
0.024417 | -34.4618 | 22.0518
0.030521 | -34.3517 | 18.2517
0.036626 | -34.2109 | 15.1623
0.04273 | -34.1284 | 12.8784
0.048834 | -34.0756 | 11.2919
0.054938 | -34.0075 | 10.0811
0.061043 | -33.934 | 8.9507
0.067147 | -33.9235 | 8.0794
0.073251 | -33.9239 | 7.3292
0.079355 | -33.9401 | 6.7808
0.08546 | -33.9565 | 6.3349
0.091564 | -33.9524 | 5.9606
0.097668 | -33.9178 | 5.6918
0.10377 | -33.8699 | 5.333
0.10988 | -33.8392 | 4.9613
0.11598 | -33.8308 | 4.5963
0.12209 | -33.8486 | 4.317
0.12819 | -33.8823 | 4.0884
0.13429 | -33.9148 | 3.8544
0.1404 | -33.9325 | 3.5875
0.1465 | -33.9489 | 3.3342
0.15261 | -33.9569 | 3.097
0.15871 | -33.9543 | 2.8986
0.16482 | -33.9402 | 2.7322
0.17092 | -33.9143 | 2.6602
0.17702 | -33.8944 | 2.5557
0.18313 | -33.8759 | 2.4686
0.18923 | -33.8625 | 2.3624
0.19534 | -33.8551 | 2.2627
0.20144 | -33.8541 | 2.1556
0.20754 | -33.8575 | 2.0873
0.21365 | -33.8601 | 2.0169
0.21975 | -33.8648 | 1.9545
0.22586 | -33.8731 | 1.8896
0.23196 | -33.8818 | 1.8142
0.23807 | -33.8904 | 1.7321
0.24417 | -33.8922 | 1.6339
0.25027 | -33.8952 | 1.5459
0.25638 | -33.8949 | 1.4501
0.26248 | -33.8927 | 1.3608
0.26859 | -33.8901 | 1.2756
0.27469 | -33.8894 | 1.188
0.2808 | -33.8955 | 1.1244
0.2869 | -33.9002 | 1.056
0.293 | -33.9059 | 0.99885
0.29911 | -33.9124 | 0.96068
0.30521 | -33.9186 | 0.92451
0.31132 | -33.9227 | 0.90986
0.31742 | -33.9271 | 0.88689
0.32353 | -33.9307 | 0.8673
0.32963 | -33.9336 | 0.85035
0.33573 | -33.9355 | 0.83459
0.34184 | -33.938 | 0.81569
0.34794 | -33.941 | 0.77805
0.35405 | -33.9442 | 0.74558
0.36015 | -33.9475 | 0.70982
0.36626 | -33.9497 | 0.67077
0.37236 | -33.9519 | 0.61963
0.37846 | -33.9538 | 0.56182
0.38457 | -33.9526 | 0.50369
0.39067 | -33.953 | 0.45344
0.39678 | -33.9511 | 0.39743
0.40288 | -33.9485 | 0.34503
0.40899 | -33.9444 | 0.29358
0.41509 | -33.9397 | 0.24157
0.42119 | -33.9343 | 0.19346
0.4273 | -33.9303 | 0.14418
0.4334 | -33.9267 | 0.094054
0.43951 | -33.9233 | 0.044431
0.44561 | -33.9216 | -0.002384
0.45172 | -33.9229 | -0.041832
0.45782 | -33.9225 | -0.085105
0.46392 | -33.9237 | -0.12483
0.47003 | -33.9253 | -0.16132
0.47613 | -33.9273 | -0.1943
0.48224 | -33.9296 | -0.22549
0.48834 | -33.9324 | -0.24691
0.49445 | -33.9347 | -0.27181
0.50055 | -33.9376 | -0.2946
0.50665 | -33.9396 | -0.3126
0.51276 | -33.9411 | -0.33223
0.51886 | -33.9427 | -0.34916
0.52497 | -33.942 | -0.35585
0.53107 | -33.9427 | -0.36574
0.53717 | -33.9434 | -0.37688
0.54328 | -33.9434 | -0.38229
0.54938 | -33.9436 | -0.39087
0.55549 | -33.9432 | -0.40055
0.56159 | -33.943 | -0.41382
0.5677 | -33.9427 | -0.42604
0.5738 | -33.9421 | -0.44098
0.5799 | -33.9421 | -0.4549
0.58601 | -33.9425 | -0.47428
0.59211 | -33.9434 | -0.50068
0.59822 | -33.944 | -0.52037
0.60432 | -33.9449 | -0.54484
0.61043 | -33.9459 | -0.56969
0.61653 | -33.947 | -0.59444
0.62263 | -33.948 | -0.621
0.62874 | -33.9494 | -0.6444
0.63484 | -33.9505 | -0.66889
0.64095 | -33.9518 | -0.69305
0.64705 | -33.953 | -0.71615
0.65316 | -33.9541 | -0.73829
0.65926 | -33.9552 | -0.7631
0.66536 | -33.9548 | -0.78296
0.67147 | -33.9552 | -0.80141
0.67757 | -33.9558 | -0.82139
0.68368 | -33.9558 | -0.84283
0.68978 | -33.9561 | -0.86236
0.69589 | -33.9558 | -0.8818
0.70199 | -33.9551 | -0.90431
0.70809 | -33.9546 | -0.92573
0.7142 | -33.9543 | -0.94595
0.7203 | -33.9535 | -0.96668
0.72641 | -33.9527 | -0.98819
0.73251 | -33.9526 | -1.0101
0.73862 | -33.9525 | -1.0355
0.74472 | -33.9529 | -1.0558
0.75082 | -33.9524 | -1.0776
0.75693 | -33.9525 | -1.1021
0.76303 | -33.9523 | -1.1255
0.76914 | -33.9526 | -1.1541
0.77524 | -33.9526 | -1.1806
0.78135 | -33.9529 | -1.2054
0.78745 | -33.9533 | -1.2317
0.79355 | -33.9536 | -1.2579
0.79966 | -33.9544 | -1.2854
0.80576 | -33.9553 | -1.3179
0.81187 | -33.9554 | -1.349
0.81797 | -33.9558 | -1.3815
0.82408 | -33.9562 | -1.4104
0.83018 | -33.9566 | -1.4435
0.83628 | -33.9572 | -1.4762
0.84239 | -33.9569 | -1.5059
0.84849 | -33.9567 | -1.5354
0.8546 | -33.957 | -1.5648
0.8607 | -33.9568 | -1.5958
0.86681 | -33.9568 | -1.6261
0.87291 | -33.9562 | -1.6535
0.87901 | -33.9556 | -1.6775
0.88512 | -33.9554 | -1.7021
0.89122 | -33.9552 | -1.7271
0.89733 | -33.9549 | -1.7481
0.90343 | -33.9545 | -1.771
0.90953 | -33.9547 | -1.7875
0.91564 | -33.9542 | -1.8049
0.92174 | -33.9538 | -1.8206
0.92785 | -33.9533 | -1.8356
0.93395 | -33.9533 | -1.8488
0.94006 | -33.9533 | -1.8642
0.94616 | -33.9532 | -1.8701
0.95226 | -33.9533 | -1.8831
0.95837 | -33.9535 | -1.8962
0.96447 | -33.9539 | -1.9079
0.97058 | -33.954 | -1.9192
0.97668 | -33.9538 | -1.9283
0.98279 | -33.9539 | -1.9328
0.98889 | -33.9543 | -1.939
0.99499 | -33.9544 | -1.9487
1.0011 | -33.9547 | -1.9549
1.0072 | -33.9547 | -1.9596
1.0133 | -33.9547 | -1.9647
1.0194 | -33.9548 | -1.9711
1.0255 | -33.9546 | -1.9788
1.0316 | -33.9541 | -1.9839
1.0377 | -33.9536 | -1.9907
1.0438 | -33.953 | -1.996
1.0499 | -33.9515 | -2.003
1.056 | -33.9504 | -2.0084
1.0621 | -33.9493 | -2.0149
1.0682 | -33.9489 | -2.0208
1.0743 | -33.9478 | -2.0293
1.0805 | -33.9473 | -2.0369
1.0866 | -33.9457 | -2.0485
1.0927 | -33.9444 | -2.0568
1.0988 | -33.9437 | -2.0684
1.1049 | -33.9423 | -2.0768
1.111 | -33.9407 | -2.0888
1.1171 | -33.9394 | -2.099
1.1232 | -33.9381 | -2.1143
1.1293 | -33.9369 | -2.1274
1.1354 | -33.9359 | -2.1408
1.1415 | -33.9347 | -2.156
1.1476 | -33.9334 | -2.171
1.1537 | -33.932 | -2.1853
1.1598 | -33.9319 | -2.2051
1.1659 | -33.9314 | -2.2184
1.172 | -33.9306 | -2.2329
1.1781 | -33.9299 | -2.2482
1.1842 | -33.9293 | -2.2652
1.1903 | -33.9301 | -2.2822
1.1964 | -33.93 | -2.2981
1.2025 | -33.9299 | -2.3148
1.2086 | -33.9304 | -2.3322
1.2147 | -33.9307 | -2.3488
1.2209 | -33.9314 | -2.3639
1.227 | -33.9333 | -2.3808
1.2331 | -33.9344 | -2.3967
1.2392 | -33.9354 | -2.4143
1.2453 | -33.9366 | -2.4279
1.2514 | -33.938 | -2.4419
1.2575 | -33.9396 | -2.4577
1.2636 | -33.942 | -2.4711
1.2697 | -33.9438 | -2.4846
1.2758 | -33.9455 | -2.5005
1.2819 | -33.9473 | -2.516
1.288 | -33.949 | -2.5315
1.2941 | -33.9512 | -2.5441
1.3002 | -33.9532 | -2.5542
1.3063 | -33.955 | -2.5659
1.3124 | -33.9569 | -2.5777
1.3185 | -33.9589 | -2.5902
1.3246 | -33.9608 | -2.6028
1.3307 | -33.963 | -2.6141
1.3368 | -33.9649 | -2.6233
1.3429 | -33.9666 | -2.6338
1.349 | -33.9681 | -2.6438
1.3551 | -33.9699 | -2.6566
1.3613 | -33.9715 | -2.6661
1.3674 | -33.9728 | -2.676
1.3735 | -33.9742 | -2.6856
1.3796 | -33.9755 | -2.6946
1.3857 | -33.9765 | -2.703
1.3918 | -33.9776 | -2.7144
1.3979 | -33.979 | -2.7244
1.404 | -33.9797 | -2.7358
1.4101 | -33.9806 | -2.7462
1.4162 | -33.9815 | -2.7561
1.4223 | -33.9823 | -2.7678
1.4284 | -33.9831 | -2.7776
1.4345 | -33.9838 | -2.7893
1.4406 | -33.9843 | -2.8015
1.4467 | -33.985 | -2.8122
1.4528 | -33.9856 | -2.8228
1.4589 | -33.9865 | -2.8347
1.465 | -33.9867 | -2.8437
1.4711 | -33.9877 | -2.854
1.4772 | -33.9882 | -2.8673
1.4833 | -33.9888 | -2.8809
1.4894 | -33.9895 | -2.8936
1.4955 | -33.9903 | -2.9066
1.5016 | -33.9909 | -2.9199
1.5078 | -33.9914 | -2.9358
1.5139 | -33.9921 | -2.9497
1.52 | -33.9926 | -2.9643
1.5261 | -33.993 | -2.9773
1.5322 | -33.9932 | -2.991
1.5383 | -33.9935 | -3.0045
1.5444 | -33.9941 | -3.0192
1.5505 | -33.9945 | -3.0337
1.5566 | -33.995 | -3.0478
1.5627 | -33.9956 | -3.0624
1.5688 | -33.9964 | -3.0768
1.5749 | -33.9966 | -3.0904
1.581 | -33.997 | -3.1086
1.5871 | -33.9976 | -3.1255
1.5932 | -33.998 | -3.1397
1.5993 | -33.9984 | -3.1579
1.6054 | -33.9988 | -3.1743
1.6115 | -33.9991 | -3.1926
1.6176 | -33.9995 | -3.2106
1.6237 | -33.9997 | -3.2263
1.6298 | -34 | -3.2432
1.6359 | -34.0005 | -3.2592
1.642 | -34.0009 | -3.2752
1.6482 | -34.0013 | -3.2938
1.6543 | -34.0018 | -3.312
1.6604 | -34.0023 | -3.3282
1.6665 | -34.0024 | -3.3447
1.6726 | -34.0027 | -3.3625
1.6787 | -34.0033 | -3.3812
1.6848 | -34.0031 | -3.3992
1.6909 | -34.0035 | -3.417
1.697 | -34.0038 | -3.4349
1.7031 | -34.0039 | -3.4521
1.7092 | -34.0039 | -3.4686
1.7153 | -34.0042 | -3.4865
1.7214 | -34.0038 | -3.5035
1.7275 | -34.0037 | -3.5209
1.7336 | -34.0035 | -3.5387
1.7397 | -34.0035 | -3.5556
1.7458 | -34.0036 | -3.574
1.7519 | -34.0034 | -3.5918
1.758 | -34.0034 | -3.6083
1.7641 | -34.0035 | -3.6245
1.7702 | -34.0033 | -3.6418
1.7763 | -34.0034 | -3.6571
1.7824 | -34.0032 | -3.675
1.7885 | -34.003 | -3.6906
1.7947 | -34.0032 | -3.7071
1.8008 | -34.0031 | -3.7216
1.8069 | -34.003 | -3.7368
1.813 | -34.0031 | -3.7526
1.8191 | -34.0028 | -3.7684
1.8252 | -34.0026 | -3.7838
1.8313 | -34.0027 | -3.7998
1.8374 | -34.0027 | -3.8149
1.8435 | -34.0028 | -3.8286
1.8496 | -34.0025 | -3.8445
1.8557 | -34.0024 | -3.8609
1.8618 | -34.0023 | -3.8741
1.8679 | -34.0024 | -3.8886
1.874 | -34.0024 | -3.9034
1.8801 | -34.0023 | -3.9178
1.8862 | -34.0024 | -3.9321
1.8923 | -34.0023 | -3.9457
1.8984 | -34.002 | -3.9578
1.9045 | -34.0019 | -3.9703
1.9106 | -34.0017 | -3.9835
1.9167 | -34.0014 | -3.9962
1.9228 | -34.0013 | -4.0085
1.9289 | -34.001 | -4.0186
1.9351 | -34.0011 | -4.0291
1.9412 | -34.0012 | -4.0408
1.9473 | -34.0012 | -4.0543
1.9534 | -34.0011 | -4.0668
1.9595 | -34.0009 | -4.0798
1.9656 | -34.0004 | -4.0918
1.9717 | -34.0001 | -4.1037
1.9778 | -34.0001 | -4.1167
1.9839 | -34.0001 | -4.1296
1.99 | -34.0002 | -4.1408
1.9961 | -33.9999 | -4.1525
2.0022 | -33.9998 | -4.1639
2.0083 | -33.9995 | -4.1752
2.0144 | -33.9995 | -4.1864
2.0205 | -33.9995 | -4.1997
2.0266 | -33.9996 | -4.2127
2.0327 | -33.9996 | -4.2253
2.0388 | -33.9998 | -4.2407
2.0449 | -33.9999 | -4.2523
2.051 | -34 | -4.2637
2.0571 | -33.9998 | -4.2765
2.0632 | -33.9997 | -4.2884
2.0693 | -33.9997 | -4.3009
2.0754 | -33.9999 | -4.3141
2.0816 | -34.0001 | -4.3275
2.0877 | -34 | -4.3399
2.0938 | -34.0002 | -4.3537
2.0999 | -34.0002 | -4.3667
2.106 | -34.0005 | -4.38
2.1121 | -34.0005 | -4.3937
2.1182 | -34.0006 | -4.4052
2.1243 | -34.0008 | -4.4172
2.1304 | -34.0009 | -4.4288
2.1365 | -34.0013 | -4.4417
2.1426 | -34.0014 | -4.4565
2.1487 | -34.0019 | -4.4696
2.1548 | -34.0023 | -4.4841
2.1609 | -34.0029 | -4.4976
2.167 | -34.0034 | -4.5117
2.1731 | -34.0037 | -4.525
2.1792 | -34.0045 | -4.5377
2.1853 | -34.005 | -4.552
2.1914 | -34.0055 | -4.5657
2.1975 | -34.0061 | -4.5793
2.2036 | -34.0066 | -4.5939
2.2097 | -34.007 | -4.6074
2.2158 | -34.008 | -4.6191
2.222 | -34.0086 | -4.6307
2.2281 | -34.009 | -4.6429
2.2342 | -34.0095 | -4.6576
2.2403 | -34.01 | -4.6691
2.2464 | -34.0109 | -4.6817
2.2525 | -34.0113 | -4.6939
2.2586 | -34.0123 | -4.7052
2.2647 | -34.0132 | -4.7159
2.2708 | -34.0138 | -4.7291
2.2769 | -34.0147 | -4.7421
2.283 | -34.0152 | -4.7529
2.2891 | -34.0158 | -4.7647
2.2952 | -34.0165 | -4.7759
2.3013 | -34.0172 | -4.7876
2.3074 | -34.0179 | -4.7986
2.3135 | -34.0184 | -4.8114
2.3196 | -34.0191 | -4.8215
2.3257 | -34.0195 | -4.8335
2.3318 | -34.0199 | -4.8439
2.3379 | -34.0205 | -4.8568
2.344 | -34.0208 | -4.8691
2.3501 | -34.0213 | -4.8813
2.3562 | -34.0217 | -4.8911
2.3623 | -34.0221 | -4.9012
2.3685 | -34.0223 | -4.9117
2.3746 | -34.0227 | -4.9236
2.3807 | -34.0232 | -4.934
2.3868 | -34.0234 | -4.9444
2.3929 | -34.0239 | -4.9551
2.399 | -34.0242 | -4.9645
2.4051 | -34.0247 | -4.9755
2.4112 | -34.0251 | -4.9859
2.4173 | -34.0254 | -4.9963
2.4234 | -34.0255 | -5.0061
2.4295 | -34.0258 | -5.0161
2.4356 | -34.0259 | -5.0268
2.4417 | -34.026 | -5.0383
2.4478 | -34.0261 | -5.0493
2.4539 | -34.0265 | -5.0585
2.46 | -34.0262 | -5.0696
2.4661 | -34.0262 | -5.0802
2.4722 | -34.0265 | -5.0904
2.4783 | -34.0267 | -5.1002
2.4844 | -34.0267 | -5.1101
2.4905 | -34.0267 | -5.1207
2.4966 | -34.0264 | -5.1309
2.5027 | -34.0264 | -5.1404
2.5089 | -34.0263 | -5.1502
2.515 | -34.0262 | -5.1609
2.5211 | -34.0261 | -5.1705
2.5272 | -34.0258 | -5.1809
2.5333 | -34.0257 | -5.1919
2.5394 | -34.0256 | -5.2023
2.5455 | -34.0255 | -5.213
2.5516 | -34.0254 | -5.2232
2.5577 | -34.0254 | -5.2345
2.5638 | -34.0252 | -5.2462
2.5699 | -34.025 | -5.2578
2.576 | -34.0248 | -5.2696
2.5821 | -34.0246 | -5.2809
2.5882 | -34.0244 | -5.2918
2.5943 | -34.0243 | -5.303
2.6004 | -34.0241 | -5.3144
2.6065 | -34.0239 | -5.3243
2.6126 | -34.0239 | -5.3348
2.6187 | -34.0238 | -5.345
2.6248 | -34.0236 | -5.3554
2.6309 | -34.0235 | -5.3681
2.637 | -34.0235 | -5.3807
2.6431 | -34.0235 | -5.3916
2.6492 | -34.0234 | -5.403
2.6554 | -34.0234 | -5.4147
2.6615 | -34.0232 | -5.4259
2.6676 | -34.0232 | -5.4377
2.6737 | -34.0233 | -5.4499
2.6798 | -34.0233 | -5.4612
2.6859 | -34.0233 | -5.4731
2.692 | -34.0232 | -5.4843
2.6981 | -34.0232 | -5.4969
2.7042 | -34.0233 | -5.5105
2.7103 | -34.0233 | -5.522
2.7164 | -34.0235 | -5.5342
2.7225 | -34.0236 | -5.547
2.7286 | -34.0237 | -5.5596
2.7347 | -34.0238 | -5.5711
2.7408 | -34.0241 | -5.5818
2.7469 | -34.024 | -5.5946
2.753 | -34.0244 | -5.6067
2.7591 | -34.0245 | -5.6201
2.7652 | -34.0245 | -5.6331
2.7713 | -34.0246 | -5.6444
2.7774 | -34.025 | -5.6573
2.7835 | -34.0252 | -5.6695
2.7896 | -34.0253 | -5.6802
2.7958 | -34.0257 | -5.692
2.8019 | -34.0263 | -5.7032
2.808 | -34.0266 | -5.7167
2.8141 | -34.027 | -5.7312
2.8202 | -34.0275 | -5.7449
2.8263 | -34.0279 | -5.7573
2.8324 | -34.0282 | -5.7712
2.8385 | -34.0286 | -5.7848
2.8446 | -34.0292 | -5.7981
2.8507 | -34.0296 | -5.8117
2.8568 | -34.0301 | -5.8244
2.8629 | -34.0306 | -5.8373
2.869 | -34.031 | -5.8501
2.8751 | -34.0315 | -5.8637
2.8812 | -34.0322 | -5.8785
2.8873 | -34.0325 | -5.8913
2.8934 | -34.0329 | -5.9047
2.8995 | -34.0333 | -5.9182
2.9056 | -34.0337 | -5.9304
2.9117 | -34.0342 | -5.9435
2.9178 | -34.0349 | -5.9578
2.9239 | -34.0353 | -5.9726
2.93 | -34.0358 | -5.9877
2.9361 | -34.0364 | -6.0006
2.9423 | -34.0371 | -6.0141
2.9484 | -34.0376 | -6.0292
2.9545 | -34.0383 | -6.0441
2.9606 | -34.0389 | -6.059
2.9667 | -34.0394 | -6.0723
2.9728 | -34.0398 | -6.0861
2.9789 | -34.0403 | -6.1003
2.985 | -34.0408 | -6.1157
2.9911 | -34.0413 | -6.1298
2.9972 | -34.0419 | -6.144
3.0033 | -34.0425 | -6.1586
3.0094 | -34.0431 | -6.1732
3.0155 | -34.0437 | -6.1874
3.0216 | -34.0443 | -6.2023
3.0277 | -34.0447 | -6.2176
3.0338 | -34.0454 | -6.2328
3.0399 | -34.046 | -6.2481
3.046 | -34.0464 | -6.2626
3.0521 | -34.047 | -6.2766
3.0582 | -34.0473 | -6.2919
3.0643 | -34.0479 | -6.3065
3.0704 | -34.0482 | -6.3214
3.0765 | -34.0486 | -6.3352
3.0827 | -34.049 | -6.3498
3.0888 | -34.0493 | -6.3642
3.0949 | -34.0497 | -6.3787
3.101 | -34.0502 | -6.3944
3.1071 | -34.0505 | -6.4101
3.1132 | -34.051 | -6.4256
3.1193 | -34.0515 | -6.4402
3.1254 | -34.0519 | -6.4538
3.1315 | -34.0524 | -6.4697
3.1376 | -34.0526 | -6.4841
3.1437 | -34.053 | -6.4992
3.1498 | -34.0534 | -6.5149
3.1559 | -34.0538 | -6.5296
3.162 | -34.0538 | -6.5452
3.1681 | -34.0539 | -6.561
3.1742 | -34.0544 | -6.5762
3.1803 | -34.0548 | -6.5909
3.1864 | -34.0552 | -6.6047
3.1925 | -34.0555 | -6.6199
3.1986 | -34.0556 | -6.6341
3.2047 | -34.0558 | -6.6497
3.2108 | -34.0561 | -6.6628
3.2169 | -34.0565 | -6.6781
3.223 | -34.0565 | -6.693
3.2292 | -34.0565 | -6.7085
3.2353 | -34.0564 | -6.7194
3.2414 | -34.0566 | -6.7339
3.2475 | -34.0566 | -6.7463
3.2536 | -34.057 | -6.7616
3.2597 | -34.0572 | -6.7743
3.2658 | -34.0574 | -6.7865
3.2719 | -34.0576 | -6.799
3.278 | -34.0577 | -6.8126
3.2841 | -34.0578 | -6.826
3.2902 | -34.058 | -6.84
3.2963 | -34.0581 | -6.8531
3.3024 | -34.058 | -6.8647
3.3085 | -34.0583 | -6.8787
3.3146 | -34.0583 | -6.8911
3.3207 | -34.0585 | -6.9047
3.3268 | -34.0587 | -6.9184
3.3329 | -34.0588 | -6.9296
3.339 | -34.0589 | -6.9406
3.3451 | -34.059 | -6.9535
3.3512 | -34.0591 | -6.9643
3.3573 | -34.0594 | -6.9771
3.3634 | -34.0595 | -6.99
3.3696 | -34.0595 | -7.0034
3.3757 | -34.0595 | -7.0147
3.3818 | -34.0598 | -7.0279
3.3879 | -34.0598 | -7.0396
3.394 | -34.0598 | -7.0515
3.4001 | -34.0598 | -7.0638
3.4062 | -34.0599 | -7.0754
3.4123 | -34.06 | -7.0867
3.4184 | -34.0601 | -7.0991
3.4245 | -34.0602 | -7.1111
3.4306 | -34.0602 | -7.1222
3.4367 | -34.0603 | -7.1334
3.4428 | -34.0607 | -7.144
3.4489 | -34.0607 | -7.1565
3.455 | -34.0606 | -7.1684
3.4611 | -34.0607 | -7.181
3.4672 | -34.0607 | -7.1938
3.4733 | -34.0608 | -7.2053
3.4794 | -34.0608 | -7.2172
3.4855 | -34.0611 | -7.2282
3.4916 | -34.0613 | -7.2394
3.4977 | -34.0612 | -7.2518
3.5038 | -34.0614 | -7.2631
3.5099 | -34.0616 | -7.2742
3.5161 | -34.0617 | -7.2852
3.5222 | -34.0619 | -7.2967
3.5283 | -34.0618 | -7.3082
3.5344 | -34.0619 | -7.3204
3.5405 | -34.0621 | -7.3317
3.5466 | -34.0623 | -7.3429
3.5527 | -34.0624 | -7.355
3.5588 | -34.0625 | -7.3661
3.5649 | -34.0626 | -7.3777
3.571 | -34.0627 | -7.3894
3.5771 | -34.0629 | -7.4015
3.5832 | -34.063 | -7.4127
3.5893 | -34.0633 | -7.4239
3.5954 | -34.0634 | -7.4359
3.6015 | -34.0636 | -7.4476
3.6076 | -34.0639 | -7.4587
3.6137 | -34.064 | -7.47
3.6198 | -34.0644 | -7.4804
3.6259 | -34.0645 | -7.4926
3.632 | -34.0648 | -7.5047
3.6381 | -34.065 | -7.5159
3.6442 | -34.0653 | -7.5278
3.6503 | -34.0655 | -7.5383
3.6565 | -34.0659 | -7.5515
3.6626 | -34.066 | -7.5637
3.6687 | -34.0663 | -7.5757
3.6748 | -34.0667 | -7.5871
3.6809 | -34.0671 | -7.599
3.687 | -34.0673 | -7.6116
3.6931 | -34.0678 | -7.623
3.6992 | -34.068 | -7.6351
3.7053 | -34.0683 | -7.6466
3.7114 | -34.0685 | -7.6594
3.7175 | -34.0689 | -7.6706
3.7236 | -34.0691 | -7.6808
3.7297 | -34.0693 | -7.693
3.7358 | -34.0696 | -7.7061
3.7419 | -34.0698 | -7.7192
3.748 | -34.07 | -7.7314
3.7541 | -34.0702 | -7.7438
3.7602 | -34.0706 | -7.7561
3.7663 | -34.0708 | -7.7677
3.7724 | -34.0712 | -7.7802
3.7785 | -34.0716 | -7.792
3.7846 | -34.0718 | -7.8041
3.7907 | -34.0723 | -7.8165
3.7969 | -34.0725 | -7.8295
3.803 | -34.0729 | -7.841
3.8091 | -34.0733 | -7.8527
3.8152 | -34.0736 | -7.8654
3.8213 | -34.0738 | -7.8784
3.8274 | -34.0743 | -7.8904
3.8335 | -34.0747 | -7.9034
3.8396 | -34.0752 | -7.9163
3.8457 | -34.0756 | -7.928
3.8518 | -34.0761 | -7.9403
3.8579 | -34.0765 | -7.9523
3.864 | -34.0769 | -7.9644
3.8701 | -34.0774 | -7.9768
3.8762 | -34.0777 | -7.9894
3.8823 | -34.078 | -8.0014
3.8884 | -34.0785 | -8.0132
3.8945 | -34.079 | -8.0259
3.9006 | -34.0795 | -8.0389
3.9067 | -34.0798 | -8.0522
3.9128 | -34.0803 | -8.0649
3.9189 | -34.0807 | -8.077
3.925 | -34.081 | -8.0892
3.9311 | -34.0815 | -8.1007
3.9372 | -34.0819 | -8.1137
3.9434 | -34.0823 | -8.1255
3.9495 | -34.0827 | -8.1385
3.9556 | -34.0832 | -8.1512
3.9617 | -34.0836 | -8.1636
3.9678 | -34.084 | -8.1757
3.9739 | -34.0845 | -8.1883
3.98 | -34.0849 | -8.2016
3.9861 | -34.0852 | -8.2135
3.9922 | -34.0855 | -8.2266
3.9983 | -34.0859 | -8.2381
4.0044 | -34.0863 | -8.2492
4.0105 | -34.0867 | -8.2615
4.0166 | -34.0873 | -8.2739
4.0227 | -34.0877 | -8.2857
4.0288 | -34.088 | -8.2985
4.0349 | -34.0883 | -8.3097
4.041 | -34.0887 | -8.3221
4.0471 | -34.089 | -8.3337
4.0532 | -34.0895 | -8.3464
4.0593 | -34.0898 | -8.3594
4.0654 | -34.0902 | -8.3716
4.0715 | -34.0906 | -8.3844
4.0776 | -34.091 | -8.3961
4.0838 | -34.0914 | -8.4089
4.0899 | -34.0918 | -8.4216
4.096 | -34.0921 | -8.4335
4.1021 | -34.0925 | -8.4457
4.1082 | -34.0929 | -8.4574
4.1143 | -34.0933 | -8.4688
4.1204 | -34.0938 | -8.4809
4.1265 | -34.0942 | -8.4928
4.1326 | -34.0946 | -8.5045
4.1387 | -34.0951 | -8.5172
4.1448 | -34.0955 | -8.5296
4.1509 | -34.0958 | -8.5418
4.157 | -34.0961 | -8.5544
4.1631 | -34.0964 | -8.567
4.1692 | -34.0969 | -8.5791
4.1753 | -34.0973 | -8.5915
4.1814 | -34.0976 | -8.6027
4.1875 | -34.0979 | -8.6155
4.1936 | -34.0983 | -8.6283
4.1997 | -34.0986 | -8.6399
4.2058 | -34.0989 | -8.6529
4.2119 | -34.0993 | -8.664
4.218 | -34.0996 | -8.6755
4.2241 | -34.1 | -8.6865
4.2303 | -34.1005 | -8.6993
4.2364 | -34.1009 | -8.7121
4.2425 | -34.1012 | -8.7241
4.2486 | -34.1015 | -8.7367
4.2547 | -34.1017 | -8.7473
4.2608 | -34.102 | -8.7586
4.2669 | -34.1023 | -8.7704
4.273 | -34.1027 | -8.7826
4.2791 | -34.103 | -8.7947
4.2852 | -34.1032 | -8.8064
4.2913 | -34.1034 | -8.817
4.2974 | -34.1037 | -8.8294
4.3035 | -34.104 | -8.8416
4.3096 | -34.1041 | -8.8534
4.3157 | -34.1045 | -8.8648
4.3218 | -34.1047 | -8.8762
4.3279 | -34.105 | -8.8879
4.334 | -34.1053 | -8.8992
4.3401 | -34.1056 | -8.9105
4.3462 | -34.1059 | -8.9228
4.3523 | -34.1063 | -8.9339
4.3584 | -34.1065 | -8.9438
4.3645 | -34.1067 | -8.9548
4.3707 | -34.107 | -8.9655
4.3768 | -34.1073 | -8.9773
4.3829 | -34.1074 | -8.989
4.389 | -34.1078 | -9.0001
4.3951 | -34.1081 | -9.011
4.4012 | -34.1083 | -9.0222
4.4073 | -34.1086 | -9.0338
4.4134 | -34.1089 | -9.0454
4.4195 | -34.1092 | -9.0564
4.4256 | -34.1095 | -9.0674
4.4317 | -34.1098 | -9.0782
4.4378 | -34.11 | -9.0893
4.4439 | -34.1103 | -9.1001
4.45 | -34.1104 | -9.1115
4.4561 | -34.1106 | -9.1227
4.4622 | -34.1108 | -9.1337
4.4683 | -34.111 | -9.1443
4.4744 | -34.1113 | -9.1559
4.4805 | -34.1116 | -9.1663
4.4866 | -34.1119 | -9.177
4.4927 | -34.1121 | -9.1886
4.4988 | -34.1123 | -9.1994
4.5049 | -34.1126 | -9.2099
4.511 | -34.1129 | -9.2203
4.5172 | -34.1133 | -9.2316
4.5233 | -34.1135 | -9.2435
4.5294 | -34.1138 | -9.2543
4.5355 | -34.1141 | -9.2651
4.5416 | -34.1144 | -9.2755
4.5477 | -34.1146 | -9.2861
4.5538 | -34.1149 | -9.2963
4.5599 | -34.1153 | -9.3072
4.566 | -34.1155 | -9.3183
4.5721 | -34.1158 | -9.3289
4.5782 | -34.1161 | -9.3401
4.5843 | -34.1164 | -9.351
4.5904 | -34.1168 | -9.3617
4.5965 | -34.1171 | -9.3718
4.6026 | -34.1174 | -9.3838
4.6087 | -34.1177 | -9.3942
4.6148 | -34.118 | -9.4062
4.6209 | -34.1184 | -9.416
4.627 | -34.1188 | -9.4275
4.6331 | -34.1189 | -9.4377
4.6392 | -34.1192 | -9.4483
4.6453 | -34.1195 | -9.4591
4.6514 | -34.1199 | -9.47
4.6576 | -34.1199 | -9.4819
4.6637 | -34.1202 | -9.4908
4.6698 | -34.1205 | -9.5007
4.6759 | -34.1207 | -9.5119
4.682 | -34.1209 | -9.5223
4.6881 | -34.1212 | -9.5325
4.6942 | -34.1215 | -9.5431
4.7003 | -34.1218 | -9.554
4.7064 | -34.1221 | -9.5646
4.7125 | -34.1224 | -9.5754
4.7186 | -34.1226 | -9.5852
4.7247 | -34.123 | -9.5961
4.7308 | -34.1234 | -9.6068
4.7369 | -34.1236 | -9.6178
4.743 | -34.1239 | -9.6295
4.7491 | -34.1241 | -9.6404
4.7552 | -34.1243 | -9.651
4.7613 | -34.1246 | -9.662
4.7674 | -34.1249 | -9.6719
4.7735 | -34.1251 | -9.6833
4.7796 | -34.1254 | -9.6944
4.7857 | -34.1255 | -9.7055
4.7918 | -34.1258 | -9.7154
4.7979 | -34.1262 | -9.7258
4.8041 | -34.1264 | -9.7364
4.8102 | -34.1268 | -9.7469
4.8163 | -34.1271 | -9.7577
4.8224 | -34.1272 | -9.7677
4.8285 | -34.1276 | -9.7781
4.8346 | -34.1279 | -9.789
4.8407 | -34.1282 | -9.7997
4.8468 | -34.1285 | -9.8107
4.8529 | -34.1287 | -9.8218
4.859 | -34.129 | -9.8326
4.8651 | -34.1293 | -9.8434
4.8712 | -34.1296 | -9.8541
4.8773 | -34.1298 | -9.865
4.8834 | -34.1302 | -9.8761
4.8895 | -34.1305 | -9.8875
4.8956 | -34.1307 | -9.8989
4.9017 | -34.1309 | -9.9097
4.9078 | -34.1312 | -9.9209
4.9139 | -34.1315 | -9.9308
4.92 | -34.1318 | -9.9412
4.9261 | -34.1322 | -9.9513
4.9322 | -34.1326 | -9.9618
4.9383 | -34.133 | -9.9722
4.9445 | -34.1334 | -9.9837
4.9506 | -34.1336 | -9.9939
4.9567 | -34.1339 | -10.0045
4.9628 | -34.1343 | -10.0152
4.9689 | -34.1345 | -10.0264
4.975 | -34.1348 | -10.037
4.9811 | -34.1351 | -10.0482
4.9872 | -34.1354 | -10.0593
4.9933 | -34.1358 | -10.0704
4.9994 | -34.1362 | -10.0815
5.0055 | -34.1366 | -10.0925
5.0116 | -34.137 | -10.1041
5.0177 | -34.1375 | -10.1155
5.0238 | -34.1378 | -10.1273
5.0299 | -34.1384 | -10.1397
5.036 | -34.1387 | -10.1506
5.0421 | -34.1389 | -10.1616
5.0482 | -34.1393 | -10.1733
5.0543 | -34.1397 | -10.1839
5.0604 | -34.1401 | -10.1956
5.0665 | -34.1405 | -10.2066
5.0726 | -34.1409 | -10.2184
5.0787 | -34.1412 | -10.23
5.0848 | -34.1416 | -10.2414
5.091 | -34.142 | -10.2526
5.0971 | -34.1424 | -10.2641
5.1032 | -34.1428 | -10.2758
5.1093 | -34.1432 | -10.2871
5.1154 | -34.1435 | -10.2985
5.1215 | -34.1438 | -10.3099
5.1276 | -34.1442 | -10.3211
5.1337 | -34.1447 | -10.3321
5.1398 | -34.145 | -10.3433
5.1459 | -34.1454 | -10.3551
5.152 | -34.1458 | -10.3675
5.1581 | -34.1461 | -10.3796
5.1642 | -34.1464 | -10.3909
5.1703 | -34.1469 | -10.4023
5.1764 | -34.1472 | -10.4136
5.1825 | -34.1476 | -10.4253
5.1886 | -34.1479 | -10.4365
5.1947 | -34.1483 | -10.4487
5.2008 | -34.1488 | -10.4599
5.2069 | -34.1492 | -10.4724
5.213 | -34.1496 | -10.4834
5.2191 | -34.15 | -10.4951
5.2252 | -34.1503 | -10.5068
5.2314 | -34.1506 | -10.5189
5.2375 | -34.1511 | -10.5302
5.2436 | -34.1514 | -10.5417
5.2497 | -34.1519 | -10.553
5.2558 | -34.1523 | -10.5654
5.2619 | -34.1527 | -10.5776
5.268 | -34.1531 | -10.5894
5.2741 | -34.1535 | -10.602
5.2802 | -34.1539 | -10.6142
5.2863 | -34.1544 | -10.6258
5.2924 | -34.1547 | -10.6383
5.2985 | -34.1552 | -10.6511
5.3046 | -34.1557 | -10.6637
5.3107 | -34.1562 | -10.6759
5.3168 | -34.1567 | -10.6878
5.3229 | -34.1571 | -10.6989
5.329 | -34.1575 | -10.7102
5.3351 | -34.158 | -10.7218
5.3412 | -34.1585 | -10.7333
5.3473 | -34.1589 | -10.7456
5.3534 | -34.1594 | -10.7579
5.3595 | -34.1598 | -10.77
5.3656 | -34.1602 | -10.7824
5.3717 | -34.1606 | -10.7943
5.3779 | -34.161 | -10.8065
5.384 | -34.1614 | -10.8187
5.3901 | -34.1618 | -10.8307
5.3962 | -34.1622 | -10.8429
5.4023 | -34.1626 | -10.8551
5.4084 | -34.1631 | -10.8669
5.4145 | -34.1635 | -10.8793
5.4206 | -34.1638 | -10.892
5.4267 | -34.1643 | -10.9047
5.4328 | -34.1647 | -10.9162
5.4389 | -34.1652 | -10.9279
5.445 | -34.1655 | -10.9406
5.4511 | -34.166 | -10.9524
5.4572 | -34.1665 | -10.9649
5.4633 | -34.1669 | -10.9767
5.4694 | -34.1672 | -10.9885
5.4755 | -34.1676 | -11.0008
5.4816 | -34.168 | -11.0128
5.4877 | -34.1685 | -11.0249
5.4938 | -34.1689 | -11.0367
5.4999 | -34.1693 | -11.0488
5.506 | -34.1697 | -11.0609
5.5121 | -34.1701 | -11.0734
5.5183 | -34.1705 | -11.0849
5.5244 | -34.1708 | -11.0971
5.5305 | -34.1713 | -11.1094
5.5366 | -34.1717 | -11.1212
5.5427 | -34.1721 | -11.1339
5.5488 | -34.1725 | -11.1463
5.5549 | -34.1728 | -11.1584
5.561 | -34.1732 | -11.1711
5.5671 | -34.1737 | -11.1829
5.5732 | -34.1741 | -11.1952
5.5793 | -34.1745 | -11.2081
5.5854 | -34.175 | -11.2205
5.5915 | -34.1753 | -11.2328
5.5976 | -34.1757 | -11.2452
5.6037 | -34.1761 | -11.2573
5.6098 | -34.1765 | -11.2692
5.6159 | -34.1769 | -11.2816
5.622 | -34.1773 | -11.2934
5.6281 | -34.1777 | -11.3056
5.6342 | -34.1782 | -11.3179
5.6403 | -34.1785 | -11.3294
5.6464 | -34.179 | -11.3416
5.6525 | -34.1793 | -11.3534
5.6586 | -34.1798 | -11.3648
5.6648 | -34.1801 | -11.3775
5.6709 | -34.1806 | -11.39
5.677 | -34.181 | -11.4029
5.6831 | -34.1813 | -11.4145
5.6892 | -34.1818 | -11.4266
5.6953 | -34.1821 | -11.4387
5.7014 | -34.1826 | -11.4514
5.7075 | -34.1829 | -11.4637
5.7136 | -34.1833 | -11.4756
5.7197 | -34.1837 | -11.4876
5.7258 | -34.184 | -11.4996
5.7319 | -34.1845 | -11.5117
5.738 | -34.1849 | -11.5237
5.7441 | -34.1853 | -11.5358
5.7502 | -34.1858 | -11.5475
5.7563 | -34.1862 | -11.5594
5.7624 | -34.1865 | -11.5722
5.7685 | -34.1869 | -11.5841
5.7746 | -34.1872 | -11.596
5.7807 | -34.1876 | -11.6077
5.7868 | -34.1879 | -11.6197
5.7929 | -34.1883 | -11.6316
5.799 | -34.1887 | -11.6435
5.8052 | -34.189 | -11.6558
5.8113 | -34.1894 | -11.6678
5.8174 | -34.1897 | -11.68
5.8235 | -34.1901 | -11.6919
5.8296 | -34.1905 | -11.7045
5.8357 | -34.1908 | -11.7166
5.8418 | -34.1912 | -11.7288
5.8479 | -34.1915 | -11.7412
5.854 | -34.1918 | -11.7535
5.8601 | -34.1922 | -11.7653
5.8662 | -34.1926 | -11.7777
5.8723 | -34.1929 | -11.7891
5.8784 | -34.1933 | -11.8011
5.8845 | -34.1937 | -11.8133
5.8906 | -34.194 | -11.8252
5.8967 | -34.1944 | -11.8368
5.9028 | -34.1949 | -11.8489
5.9089 | -34.1952 | -11.8603
5.915 | -34.1956 | -11.8723
5.9211 | -34.196 | -11.8843
5.9272 | -34.1963 | -11.8962
5.9333 | -34.1967 | -11.9075
5.9394 | -34.1971 | -11.9193
5.9455 | -34.1975 | -11.9308
5.9517 | -34.1979 | -11.9428
5.9578 | -34.1983 | -11.955
5.9639 | -34.1987 | -11.9666
5.97 | -34.1991 | -11.9786
5.9761 | -34.1995 | -11.9902
5.9822 | -34.1999 | -12.0022
5.9883 | -34.2004 | -12.0139
5.9944 | -34.2007 | -12.0259
6.0005 | -34.2012 | -12.038
6.0066 | -34.2015 | -12.0495
6.0127 | -34.2019 | -12.0606
6.0188 | -34.2023 | -12.0729
6.0249 | -34.2027 | -12.0847
6.031 | -34.2031 | -12.0963
6.0371 | -34.2035 | -12.1084
6.0432 | -34.204 | -12.1203
6.0493 | -34.2044 | -12.1319
6.0554 | -34.2049 | -12.1436
6.0615 | -34.2052 | -12.156
6.0676 | -34.2056 | -12.1685
6.0737 | -34.206 | -12.1805
6.0798 | -34.2064 | -12.1921
6.0859 | -34.2068 | -12.2039
6.0921 | -34.2072 | -12.2155
6.0982 | -34.2077 | -12.2276
6.1043 | -34.2081 | -12.2392
6.1104 | -34.2086 | -12.2516
6.1165 | -34.209 | -12.2626
6.1226 | -34.2094 | -12.2738
6.1287 | -34.2098 | -12.2857
6.1348 | -34.2103 | -12.2976
6.1409 | -34.2107 | -12.3092
6.147 | -34.2111 | -12.3208
6.1531 | -34.2116 | -12.3326
6.1592 | -34.212 | -12.3443
6.1653 | -34.2124 | -12.3559
6.1714 | -34.2128 | -12.3676
6.1775 | -34.2132 | -12.3797
6.1836 | -34.2136 | -12.3912
6.1897 | -34.214 | -12.4029
6.1958 | -34.2144 | -12.4149
6.2019 | -34.2148 | -12.4269
6.208 | -34.2153 | -12.4391
6.2141 | -34.2157 | -12.451
6.2202 | -34.2161 | -12.4632
6.2263 | -34.2166 | -12.4748
6.2325 | -34.217 | -12.4861
6.2386 | -34.2175 | -12.4975
6.2447 | -34.2179 | -12.5094
6.2508 | -34.2183 | -12.521
6.2569 | -34.2187 | -12.5331
6.263 | -34.2192 | -12.5445
6.2691 | -34.2197 | -12.5561
6.2752 | -34.2201 | -12.5677
6.2813 | -34.2205 | -12.5794
6.2874 | -34.2209 | -12.5911
6.2935 | -34.2214 | -12.6025
6.2996 | -34.2218 | -12.6143
6.3057 | -34.2223 | -12.6259
6.3118 | -34.2227 | -12.6375
6.3179 | -34.2231 | -12.6492
6.324 | -34.2236 | -12.6611
6.3301 | -34.2241 | -12.6726
6.3362 | -34.2245 | -12.6843
6.3423 | -34.225 | -12.6961
6.3484 | -34.2255 | -12.708
6.3545 | -34.2259 | -12.7191
6.3606 | -34.2264 | -12.7305
6.3667 | -34.2269 | -12.7418
6.3728 | -34.2274 | -12.7535
6.379 | -34.2278 | -12.7652
6.3851 | -34.2282 | -12.7769
6.3912 | -34.2287 | -12.7883
6.3973 | -34.2292 | -12.7996
6.4034 | -34.2296 | -12.8112
6.4095 | -34.23 | -12.8231
6.4156 | -34.2305 | -12.8343
6.4217 | -34.2309 | -12.8459
6.4278 | -34.2313 | -12.8574
6.4339 | -34.2319 | -12.869
6.44 | -34.2323 | -12.8802
6.4461 | -34.2327 | -12.8923
6.4522 | -34.2332 | -12.9033
6.4583 | -34.2336 | -12.915
6.4644 | -34.234 | -12.9269
6.4705 | -34.2345 | -12.9387
6.4766 | -34.235 | -12.9507
6.4827 | -34.2354 | -12.9628
6.4888 | -34.2358 | -12.9743
6.4949 | -34.2363 | -12.9854
6.501 | -34.2367 | -12.9973
6.5071 | -34.2372 | -13.0091
6.5132 | -34.2377 | -13.0205
6.5194 | -34.2381 | -13.032
6.5255 | -34.2386 | -13.0432
6.5316 | -34.239 | -13.055
6.5377 | -34.2395 | -13.0668
6.5438 | -34.2399 | -13.0783
6.5499 | -34.2404 | -13.0897
6.556 | -34.2409 | -13.1011
6.5621 | -34.2413 | -13.1128
6.5682 | -34.2417 | -13.1242
6.5743 | -34.2422 | -13.1366
6.5804 | -34.2425 | -13.1481
6.5865 | -34.2429 | -13.16
6.5926 | -34.2434 | -13.1715
6.5987 | -34.2439 | -13.1828
6.6048 | -34.2443 | -13.195
6.6109 | -34.2448 | -13.2065
6.617 | -34.2452 | -13.219
6.6231 | -34.2456 | -13.2308
6.6292 | -34.2461 | -13.2423
6.6353 | -34.2466 | -13.2549
6.6414 | -34.2471 | -13.2668
6.6475 | -34.2476 | -13.2786
6.6536 | -34.248 | -13.2898
6.6597 | -34.2485 | -13.3012
6.6659 | -34.249 | -13.3126
6.672 | -34.2495 | -13.3243
6.6781 | -34.2499 | -13.3356
6.6842 | -34.2504 | -13.3474
6.6903 | -34.2508 | -13.3589
6.6964 | -34.2513 | -13.3706
6.7025 | -34.2517 | -13.3822
6.7086 | -34.2522 | -13.3935
6.7147 | -34.2526 | -13.4053
6.7208 | -34.2531 | -13.4174
6.7269 | -34.2535 | -13.4286
6.733 | -34.2538 | -13.4394
6.7391 | -34.2544 | -13.4512
6.7452 | -34.2549 | -13.4622
6.7513 | -34.2554 | -13.4735
6.7574 | -34.2559 | -13.4853
6.7635 | -34.2564 | -13.4964
6.7696 | -34.2569 | -13.508
6.7757 | -34.2574 | -13.5199
6.7818 | -34.2578 | -13.5316
6.7879 | -34.2583 | -13.5432
6.794 | -34.2588 | -13.5546
6.8001 | -34.2592 | -13.566
6.8063 | -34.2597 | -13.5777
6.8124 | -34.2601 | -13.5894
6.8185 | -34.2605 | -13.6009
6.8246 | -34.261 | -13.6125
6.8307 | -34.2614 | -13.6239
6.8368 | -34.2619 | -13.6357
6.8429 | -34.2624 | -13.6471
6.849 | -34.2628 | -13.6587
6.8551 | -34.2633 | -13.6702
6.8612 | -34.2638 | -13.6816
6.8673 | -34.2643 | -13.6929
6.8734 | -34.2647 | -13.7052
6.8795 | -34.2652 | -13.717
6.8856 | -34.2657 | -13.728
6.8917 | -34.2662 | -13.7389
6.8978 | -34.2666 | -13.7504
6.9039 | -34.267 | -13.7618
6.91 | -34.2675 | -13.7734
6.9161 | -34.268 | -13.7847
6.9222 | -34.2684 | -13.7963
6.9283 | -34.2688 | -13.8074
6.9344 | -34.2693 | -13.8187
6.9405 | -34.2697 | -13.83
6.9466 | -34.2702 | -13.8417
6.9528 | -34.2706 | -13.8532
6.9589 | -34.2711 | -13.8645
6.965 | -34.2716 | -13.8759
6.9711 | -34.272 | -13.8874
6.9772 | -34.2725 | -13.8986
6.9833 | -34.2729 | -13.9103
6.9894 | -34.2734 | -13.9222
6.9955 | -34.2738 | -13.9336
7.0016 | -34.2743 | -13.9455
7.0077 | -34.2747 | -13.9573
7.0138 | -34.2752 | -13.9691
7.0199 | -34.2756 | -13.9803
7.026 | -34.276 | -13.9917
7.0321 | -34.2764 | -14.0032
7.0382 | -34.2768 | -14.0144
7.0443 | -34.2774 | -14.0262
7.0504 | -34.2778 | -14.038
7.0565 | -34.2782 | -14.0496
7.0626 | -34.2786 | -14.061
7.0687 | -34.2791 | -14.072
7.0748 | -34.2795 | -14.0833
7.0809 | -34.28 | -14.0949
7.087 | -34.2804 | -14.1064
7.0932 | -34.2808 | -14.1178
7.0993 | -34.2812 | -14.1293
7.1054 | -34.2817 | -14.1404
7.1115 | -34.282 | -14.1518
7.1176 | -34.2824 | -14.163
7.1237 | -34.2828 | -14.1744
7.1298 | -34.2832 | -14.1854
7.1359 | -34.2836 | -14.1964
7.142 | -34.284 | -14.2079
7.1481 | -34.2844 | -14.2191
7.1542 | -34.2848 | -14.2303
7.1603 | -34.2852 | -14.2415
7.1664 | -34.2857 | -14.253
7.1725 | -34.2861 | -14.2641
7.1786 | -34.2866 | -14.2756
7.1847 | -34.287 | -14.2868
7.1908 | -34.2874 | -14.2982
7.1969 | -34.2878 | -14.3095
7.203 | -34.2883 | -14.3209
7.2091 | -34.2887 | -14.332
7.2152 | -34.2891 | -14.3433
7.2213 | -34.2895 | -14.3544
7.2274 | -34.2899 | -14.3658
7.2335 | -34.2903 | -14.3771
7.2397 | -34.2908 | -14.3884
7.2458 | -34.2911 | -14.3992
7.2519 | -34.2916 | -14.4105
7.258 | -34.292 | -14.422
7.2641 | -34.2924 | -14.4331
7.2702 | -34.2929 | -14.4447
7.2763 | -34.2933 | -14.4558
7.2824 | -34.2938 | -14.4669
7.2885 | -34.2942 | -14.4781
7.2946 | -34.2947 | -14.4896
7.3007 | -34.295 | -14.5009
7.3068 | -34.2954 | -14.5123
7.3129 | -34.2958 | -14.5233
7.319 | -34.2963 | -14.5347
7.3251 | -34.2967 | -14.5456
7.3312 | -34.2971 | -14.5565
7.3373 | -34.2976 | -14.5675
7.3434 | -34.2981 | -14.5784
7.3495 | -34.2985 | -14.5891
7.3556 | -34.2989 | -14.6008
7.3617 | -34.2993 | -14.6125
7.3678 | -34.2998 | -14.6238
7.3739 | -34.3002 | -14.6346
7.3801 | -34.3006 | -14.6456
7.3862 | -34.301 | -14.6567
7.3923 | -34.3014 | -14.6679
7.3984 | -34.3019 | -14.6789
7.4045 | -34.3023 | -14.69
7.4106 | -34.3027 | -14.701
7.4167 | -34.3031 | -14.7124
7.4228 | -34.3036 | -14.7235
7.4289 | -34.304 | -14.7343
7.435 | -34.3045 | -14.7452
7.4411 | -34.3049 | -14.756
7.4472 | -34.3053 | -14.7673
7.4533 | -34.3058 | -14.7784
7.4594 | -34.3062 | -14.7899
7.4655 | -34.3067 | -14.8009
7.4716 | -34.307 | -14.8123
7.4777 | -34.3074 | -14.8229
7.4838 | -34.3079 | -14.8342
7.4899 | -34.3084 | -14.8453
7.496 | -34.3088 | -14.8565
7.5021 | -34.3092 | -14.8675
7.5082 | -34.3096 | -14.8787
7.5143 | -34.3101 | -14.8897
7.5204 | -34.3105 | -14.9004
7.5266 | -34.3108 | -14.9116
7.5327 | -34.3113 | -14.9225
7.5388 | -34.3117 | -14.9335
7.5449 | -34.312 | -14.9448
7.551 | -34.3125 | -14.9561
7.5571 | -34.3129 | -14.9671
7.5632 | -34.3133 | -14.9781
7.5693 | -34.3138 | -14.9893
7.5754 | -34.3142 | -15.0004
7.5815 | -34.3147 | -15.0111
7.5876 | -34.315 | -15.0221
7.5937 | -34.3154 | -15.033
7.5998 | -34.3158 | -15.0442
7.6059 | -34.3162 | -15.0552
7.612 | -34.3167 | -15.0661
7.6181 | -34.3171 | -15.0771
7.6242 | -34.3175 | -15.0882
7.6303 | -34.318 | -15.099
7.6364 | -34.3184 | -15.1103
7.6425 | -34.3188 | -15.1216
7.6486 | -34.3192 | -15.1327
7.6547 | -34.3197 | -15.1439
7.6608 | -34.3201 | -15.1549
7.667 | -34.3205 | -15.1653
7.6731 | -34.3209 | -15.1765
7.6792 | -34.3213 | -15.1877
7.6853 | -34.3216 | -15.1986
7.6914 | -34.322 | -15.2098
7.6975 | -34.3224 | -15.221
7.7036 | -34.3228 | -15.2322
7.7097 | -34.3232 | -15.2431
7.7158 | -34.3237 | -15.254
7.7219 | -34.324 | -15.2649
7.728 | -34.3244 | -15.2761
7.7341 | -34.3248 | -15.2872
7.7402 | -34.3253 | -15.2981
7.7463 | -34.3256 | -15.3092
7.7524 | -34.326 | -15.3205
7.7585 | -34.3264 | -15.3317
7.7646 | -34.3268 | -15.3424
7.7707 | -34.3271 | -15.3532
7.7768 | -34.3275 | -15.364
7.7829 | -34.3278 | -15.3752
7.789 | -34.3282 | -15.3859
7.7951 | -34.3286 | -15.3973
7.8012 | -34.329 | -15.4087
7.8073 | -34.3294 | -15.4195
7.8135 | -34.3298 | -15.4307
7.8196 | -34.3302 | -15.4416
7.8257 | -34.3306 | -15.453
7.8318 | -34.331 | -15.4645
7.8379 | -34.3314 | -15.4751
7.844 | -34.3318 | -15.4856
7.8501 | -34.3321 | -15.4969
7.8562 | -34.3324 | -15.5086
7.8623 | -34.3329 | -15.5197
7.8684 | -34.3332 | -15.5304
7.8745 | -34.3336 | -15.5413
7.8806 | -34.334 | -15.5517
7.8867 | -34.3344 | -15.5626
7.8928 | -34.3348 | -15.5735
7.8989 | -34.3353 | -15.584
7.905 | -34.3357 | -15.5954
7.9111 | -34.3361 | -15.6064
7.9172 | -34.3365 | -15.6172
7.9233 | -34.3368 | -15.6282
7.9294 | -34.3373 | -15.6393
7.9355 | -34.3376 | -15.6498
7.9416 | -34.338 | -15.6607
7.9477 | -34.3384 | -15.6712
7.9539 | -34.3388 | -15.6818
7.96 | -34.3392 | -15.693
7.9661 | -34.3395 | -15.7041
7.9722 | -34.34 | -15.7151
7.9783 | -34.3403 | -15.7261
7.9844 | -34.3408 | -15.7374
7.9905 | -34.3412 | -15.7482
7.9966 | -34.3415 | -15.7592
8.0027 | -34.3419 | -15.7705
8.0088 | -34.3423 | -15.7811
8.0149 | -34.3426 | -15.7918
8.021 | -34.343 | -15.8027
8.0271 | -34.3434 | -15.8141
8.0332 | -34.3438 | -15.8251
8.0393 | -34.3442 | -15.8362
8.0454 | -34.3446 | -15.8474
8.0515 | -34.345 | -15.8582
8.0576 | -34.3453 | -15.8695
8.0637 | -34.3457 | -15.8799
8.0698 | -34.3461 | -15.8905
8.0759 | -34.3465 | -15.9011
8.082 | -34.3469 | -15.9123
8.0881 | -34.3472 | -15.923
8.0942 | -34.3476 | -15.9345
8.1004 | -34.348 | -15.9458
8.1065 | -34.3484 | -15.9567
8.1126 | -34.3488 | -15.9675
8.1187 | -34.3492 | -15.9786
8.1248 | -34.3496 | -15.9898
8.1309 | -34.35 | -16.0007
8.137 | -34.3503 | -16.0116
8.1431 | -34.3508 | -16.0227
8.1492 | -34.3512 | -16.0338
8.1553 | -34.3516 | -16.0447
8.1614 | -34.3519 | -16.0559
8.1675 | -34.3523 | -16.0667
8.1736 | -34.3527 | -16.078
8.1797 | -34.353 | -16.0891
8.1858 | -34.3534 | -16.1
8.1919 | -34.3538 | -16.1108
8.198 | -34.3541 | -16.1214
8.2041 | -34.3546 | -16.1327
8.2102 | -34.355 | -16.1437
8.2163 | -34.3554 | -16.1546
8.2224 | -34.3558 | -16.1657
8.2285 | -34.3562 | -16.1765
8.2346 | -34.3566 | -16.1874
8.2408 | -34.357 | -16.1985
8.2469 | -34.3574 | -16.2093
8.253 | -34.3577 | -16.2201
8.2591 | -34.3581 | -16.2308
8.2652 | -34.3586 | -16.2415
8.2713 | -34.359 | -16.2525
8.2774 | -34.3594 | -16.2635
8.2835 | -34.3598 | -16.2749
8.2896 | -34.3602 | -16.2853
8.2957 | -34.3606 | -16.2962
8.3018 | -34.361 | -16.3071
8.3079 | -34.3614 | -16.318
8.314 | -34.3618 | -16.3287
8.3201 | -34.3622 | -16.3396
8.3262 | -34.3626 | -16.3505
8.3323 | -34.363 | -16.3613
8.3384 | -34.3634 | -16.3726
8.3445 | -34.3638 | -16.3831
8.3506 | -34.3642 | -16.3943
8.3567 | -34.3646 | -16.4054
8.3628 | -34.3649 | -16.4163
8.3689 | -34.3653 | -16.4272
8.375 | -34.3657 | -16.4382
8.3812 | -34.3661 | -16.4491
8.3873 | -34.3665 | -16.4601
8.3934 | -34.3669 | -16.471
8.3995 | -34.3673 | -16.482
8.4056 | -34.3677 | -16.4931
8.4117 | -34.3681 | -16.5043
8.4178 | -34.3685 | -16.5148
8.4239 | -34.3689 | -16.5254
8.43 | -34.3694 | -16.5364
8.4361 | -34.3698 | -16.5477
8.4422 | -34.3701 | -16.5585
8.4483 | -34.3705 | -16.5694
8.4544 | -34.3709 | -16.5806
8.4605 | -34.3713 | -16.5915
8.4666 | -34.3716 | -16.6026
8.4727 | -34.372 | -16.6134
8.4788 | -34.3724 | -16.6241
8.4849 | -34.3728 | -16.6354
8.491 | -34.3732 | -16.6461
8.4971 | -34.3736 | -16.6571
8.5032 | -34.374 | -16.6681
8.5093 | -34.3744 | -16.679
8.5154 | -34.3748 | -16.6896
8.5215 | -34.3751 | -16.7006
8.5277 | -34.3755 | -16.7116
8.5338 | -34.3759 | -16.7222
8.5399 | -34.3763 | -16.7334
8.546 | -34.3767 | -16.744
8.5521 | -34.3771 | -16.755
8.5582 | -34.3775 | -16.7661
8.5643 | -34.3779 | -16.7769
8.5704 | -34.3783 | -16.7875
8.5765 | -34.3787 | -16.7983
8.5826 | -34.3792 | -16.8092
8.5887 | -34.3796 | -16.8201
8.5948 | -34.3799 | -16.8306
8.6009 | -34.3804 | -16.8415
8.607 | -34.3807 | -16.8526
8.6131 | -34.3811 | -16.8638
8.6192 | -34.3815 | -16.8743
8.6253 | -34.3819 | -16.8853
8.6314 | -34.3823 | -16.8962
8.6375 | -34.3827 | -16.9073
8.6436 | -34.3832 | -16.9183
8.6497 | -34.3836 | -16.9292
8.6558 | -34.384 | -16.9399
8.6619 | -34.3844 | -16.951
8.6681 | -34.3847 | -16.9618
8.6742 | -34.3852 | -16.973
8.6803 | -34.3856 | -16.9841
8.6864 | -34.3861 | -16.9953
8.6925 | -34.3865 | -17.0062
8.6986 | -34.3869 | -17.0175
8.7047 | -34.3873 | -17.0284
8.7108 | -34.3877 | -17.0395
8.7169 | -34.3881 | -17.0501
8.723 | -34.3886 | -17.061
8.7291 | -34.3889 | -17.0718
8.7352 | -34.3894 | -17.083
8.7413 | -34.3898 | -17.0938
8.7474 | -34.3902 | -17.1044
8.7535 | -34.3907 | -17.1155
8.7596 | -34.3911 | -17.1269
8.7657 | -34.3915 | -17.1378
8.7718 | -34.3919 | -17.1485
8.7779 | -34.3923 | -17.16
8.784 | -34.3927 | -17.1707
8.7901 | -34.393 | -17.1821
8.7962 | -34.3935 | -17.1931
8.8023 | -34.394 | -17.2037
8.8084 | -34.3944 | -17.2148
8.8146 | -34.3948 | -17.2256
8.8207 | -34.3952 | -17.2366
8.8268 | -34.3957 | -17.2474
8.8329 | -34.3961 | -17.2584
8.839 | -34.3965 | -17.2694
8.8451 | -34.3969 | -17.2801
8.8512 | -34.3974 | -17.2911
8.8573 | -34.3978 | -17.3019
8.8634 | -34.3983 | -17.3127
8.8695 | -34.3987 | -17.3237
8.8756 | -34.3991 | -17.3344
8.8817 | -34.3996 | -17.3456
8.8878 | -34.3999 | -17.3568
8.8939 | -34.4003 | -17.3674
8.9 | -34.4007 | -17.378
8.9061 | -34.4012 | -17.3889
8.9122 | -34.4016 | -17.3997
8.9183 | -34.402 | -17.4105
8.9244 | -34.4024 | -17.4211
8.9305 | -34.4029 | -17.432
8.9366 | -34.4034 | -17.4428
8.9427 | -34.4038 | -17.4538
8.9488 | -34.4043 | -17.4647
8.955 | -34.4047 | -17.4756
8.9611 | -34.4051 | -17.4866
8.9672 | -34.4056 | -17.4978
8.9733 | -34.406 | -17.5086
8.9794 | -34.4065 | -17.5197
8.9855 | -34.4069 | -17.5304
8.9916 | -34.4073 | -17.5412
8.9977 | -34.4078 | -17.552
9.0038 | -34.4083 | -17.563
9.0099 | -34.4087 | -17.574
9.016 | -34.4092 | -17.5849
9.0221 | -34.4096 | -17.5961
9.0282 | -34.4101 | -17.6072
9.0343 | -34.4105 | -17.6184
9.0404 | -34.4111 | -17.6291
9.0465 | -34.4115 | -17.6404
9.0526 | -34.4119 | -17.6512
9.0587 | -34.4124 | -17.662
9.0648 | -34.4129 | -17.6729
9.0709 | -34.4133 | -17.6835
9.077 | -34.4138 | -17.695
9.0831 | -34.4142 | -17.7059
9.0892 | -34.4146 | -17.7168
9.0953 | -34.415 | -17.7278
9.1015 | -34.4155 | -17.7387
9.1076 | -34.416 | -17.7497
9.1137 | -34.4165 | -17.7605
9.1198 | -34.4169 | -17.771
9.1259 | -34.4174 | -17.782
9.132 | -34.4179 | -17.7931
9.1381 | -34.4184 | -17.8039
9.1442 | -34.4189 | -17.8144
9.1503 | -34.4193 | -17.8254
9.1564 | -34.4198 | -17.8363
9.1625 | -34.4202 | -17.8472
9.1686 | -34.4207 | -17.8581
9.1747 | -34.4211 | -17.8692
9.1808 | -34.4216 | -17.8803
9.1869 | -34.4221 | -17.8915
9.193 | -34.4225 | -17.9026
9.1991 | -34.4229 | -17.9132
9.2052 | -34.4234 | -17.924
9.2113 | -34.4239 | -17.9351
9.2174 | -34.4244 | -17.9466
9.2235 | -34.4249 | -17.9572
9.2296 | -34.4254 | -17.9683
9.2357 | -34.4258 | -17.9792
9.2419 | -34.4263 | -17.9901
9.248 | -34.4268 | -18.0012
9.2541 | -34.4273 | -18.0123
9.2602 | -34.4277 | -18.0233
9.2663 | -34.4282 | -18.0341
9.2724 | -34.4287 | -18.0451
9.2785 | -34.4291 | -18.056
9.2846 | -34.4296 | -18.067
9.2907 | -34.4301 | -18.0778
9.2968 | -34.4305 | -18.0889
9.3029 | -34.431 | -18.0999
9.309 | -34.4315 | -18.1107
9.3151 | -34.4319 | -18.1215
9.3212 | -34.4325 | -18.1326
9.3273 | -34.433 | -18.1434
9.3334 | -34.4335 | -18.1541
9.3395 | -34.434 | -18.1649
9.3456 | -34.4345 | -18.1756
9.3517 | -34.435 | -18.1861
9.3578 | -34.4355 | -18.1968
9.3639 | -34.436 | -18.2079
9.37 | -34.4365 | -18.2189
9.3761 | -34.437 | -18.2296
9.3822 | -34.4375 | -18.2404
9.3884 | -34.438 | -18.2513
9.3945 | -34.4385 | -18.2623
9.4006 | -34.4389 | -18.2732
9.4067 | -34.4394 | -18.2843
9.4128 | -34.4399 | -18.2951
9.4189 | -34.4404 | -18.3061
9.425 | -34.4409 | -18.3168
9.4311 | -34.4414 | -18.3277
9.4372 | -34.4419 | -18.3386
9.4433 | -34.4424 | -18.3496
9.4494 | -34.4428 | -18.3604
9.4555 | -34.4433 | -18.3713
9.4616 | -34.4439 | -18.382
9.4677 | -34.4444 | -18.3927
9.4738 | -34.4449 | -18.404
9.4799 | -34.4453 | -18.4148
9.486 | -34.4458 | -18.426
9.4921 | -34.4463 | -18.4373
9.4982 | -34.4468 | -18.4477
9.5043 | -34.4473 | -18.4583
9.5104 | -34.4478 | -18.469
9.5165 | -34.4483 | -18.4799
9.5226 | -34.4488 | -18.4909
9.5288 | -34.4493 | -18.5017
9.5349 | -34.4498 | -18.5126
9.541 | -34.4503 | -18.5235
9.5471 | -34.4508 | -18.5344
9.5532 | -34.4513 | -18.5452
9.5593 | -34.4519 | -18.5561
9.5654 | -34.4524 | -18.567
9.5715 | -34.4529 | -18.5781
9.5776 | -34.4534 | -18.5892
9.5837 | -34.4539 | -18.5998
9.5898 | -34.4543 | -18.6109
9.5959 | -34.4548 | -18.6216
9.602 | -34.4553 | -18.6325
9.6081 | -34.4558 | -18.6435
9.6142 | -34.4563 | -18.6543
9.6203 | -34.4568 | -18.6652
9.6264 | -34.4572 | -18.6757
9.6325 | -34.4577 | -18.6868
9.6386 | -34.4582 | -18.6975
9.6447 | -34.4587 | -18.7083
9.6508 | -34.4592 | -18.7192
9.6569 | -34.4597 | -18.73
9.663 | -34.4602 | -18.7409
9.6691 | -34.4607 | -18.7517
9.6753 | -34.4613 | -18.7626
9.6814 | -34.4618 | -18.7732
9.6875 | -34.4624 | -18.7838
9.6936 | -34.4629 | -18.7946
9.6997 | -34.4634 | -18.8056
9.7058 | -34.4639 | -18.8163
9.7119 | -34.4644 | -18.8271
9.718 | -34.4649 | -18.8382
9.7241 | -34.4653 | -18.8491
9.7302 | -34.4658 | -18.86
9.7363 | -34.4663 | -18.8707
9.7424 | -34.4669 | -18.8813
9.7485 | -34.4674 | -18.892
9.7546 | -34.4679 | -18.9026
9.7607 | -34.4684 | -18.9136
9.7668 | -34.4689 | -18.9249
9.7729 | -34.4694 | -18.9354
9.779 | -34.47 | -18.9462
9.7851 | -34.4704 | -18.9572
9.7912 | -34.471 | -18.9686
9.7973 | -34.4714 | -18.9796
9.8034 | -34.472 | -18.9907
9.8095 | -34.4724 | -19.0013
9.8157 | -34.4729 | -19.0121
9.8218 | -34.4734 | -19.0226
9.8279 | -34.474 | -19.0335
9.834 | -34.4745 | -19.0442
9.8401 | -34.475 | -19.0549
9.8462 | -34.4756 | -19.0654
9.8523 | -34.4761 | -19.0762
9.8584 | -34.4766 | -19.0873
9.8645 | -34.4771 | -19.0981
9.8706 | -34.4776 | -19.1088
9.8767 | -34.4781 | -19.1192
9.8828 | -34.4786 | -19.1296
9.8889 | -34.4792 | -19.1407
9.895 | -34.4796 | -19.1511
9.9011 | -34.4801 | -19.1614
9.9072 | -34.4806 | -19.1718
9.9133 | -34.4812 | -19.1824
9.9194 | -34.4817 | -19.1928
9.9255 | -34.4822 | -19.2038
9.9316 | -34.4828 | -19.2142
9.9377 | -34.4833 | -19.2248
9.9438 | -34.4838 | -19.2354
9.9499 | -34.4843 | -19.2462
9.956 | -34.4849 | -19.2567
9.9622 | -34.4854 | -19.2674
9.9683 | -34.4859 | -19.2781
9.9744 | -34.4864 | -19.2888
9.9805 | -34.4869 | -19.2995
9.9866 | -34.4874 | -19.3099
9.9927 | -34.488 | -19.3208
9.9988 | -34.4885 | -19.3313
10.0049 | -34.489 | -19.3419
10.011 | -34.4896 | -19.3525
10.0171 | -34.4901 | -19.3632
10.0232 | -34.4906 | -19.3736
10.0293 | -34.4912 | -19.384
10.0354 | -34.4918 | -19.3944
10.0415 | -34.4923 | -19.405
10.0476 | -34.4928 | -19.416
10.0537 | -34.4932 | -19.4269
10.0598 | -34.4938 | -19.4377
10.0659 | -34.4943 | -19.4483
10.072 | -34.4948 | -19.4588
10.0781 | -34.4954 | -19.4695
10.0842 | -34.4959 | -19.4803
10.0903 | -34.4964 | -19.4913
10.0964 | -34.497 | -19.5016
10.1026 | -34.4976 | -19.5122
10.1087 | -34.4981 | -19.5229
10.1148 | -34.4986 | -19.5336
10.1209 | -34.4992 | -19.5441
10.127 | -34.4997 | -19.5545
10.1331 | -34.5003 | -19.565
10.1392 | -34.5008 | -19.5755
10.1453 | -34.5013 | -19.5862
10.1514 | -34.5018 | -19.5965
10.1575 | -34.5023 | -19.607
10.1636 | -34.5029 | -19.6169
10.1697 | -34.5034 | -19.6274
10.1758 | -34.504 | -19.6379
10.1819 | -34.5045 | -19.6486
10.188 | -34.505 | -19.6589
10.1941 | -34.5055 | -19.6696
10.2002 | -34.506 | -19.6795
10.2063 | -34.5065 | -19.6903
10.2124 | -34.507 | -19.7009
10.2185 | -34.5076 | -19.7114
10.2246 | -34.5081 | -19.7219
10.2307 | -34.5086 | -19.7326
10.2368 | -34.5091 | -19.7423
10.2429 | -34.5097 | -19.7525
10.2491 | -34.5102 | -19.7628
10.2552 | -34.5107 | -19.7737
10.2613 | -34.5113 | -19.7845
10.2674 | -34.5118 | -19.7952
10.2735 | -34.5123 | -19.8059
10.2796 | -34.5128 | -19.8166
10.2857 | -34.5133 | -19.8271
10.2918 | -34.5138 | -19.8377
10.2979 | -34.5144 | -19.8485
10.304 | -34.515 | -19.8592
10.3101 | -34.5155 | -19.8697
10.3162 | -34.516 | -19.8802
10.3223 | -34.5165 | -19.8907
10.3284 | -34.5171 | -19.9012
10.3345 | -34.5176 | -19.9115
10.3406 | -34.5181 | -19.9218
10.3467 | -34.5186 | -19.9323
10.3528 | -34.5191 | -19.9425
10.3589 | -34.5196 | -19.9532
10.365 | -34.5202 | -19.964
10.3711 | -34.5206 | -19.9747
10.3772 | -34.5211 | -19.9852
10.3833 | -34.5216 | -19.9957
10.3895 | -34.5221 | -20.006
10.3956 | -34.5227 | -20.0163
10.4017 | -34.5232 | -20.0268
10.4078 | -34.5237 | -20.0373
10.4139 | -34.5243 | -20.0477
10.42 | -34.5248 | -20.0582
10.4261 | -34.5254 | -20.0686
10.4322 | -34.5259 | -20.0789
10.4383 | -34.5264 | -20.0894
10.4444 | -34.5269 | -20.0994
10.4505 | -34.5275 | -20.1099
10.4566 | -34.528 | -20.1203
10.4627 | -34.5285 | -20.1307
10.4688 | -34.529 | -20.1407
10.4749 | -34.5296 | -20.151
10.481 | -34.5301 | -20.1614
10.4871 | -34.5306 | -20.172
10.4932 | -34.5312 | -20.1823
10.4993 | -34.5318 | -20.1926
10.5054 | -34.5323 | -20.2031
10.5115 | -34.5328 | -20.2136
10.5176 | -34.5334 | -20.2237
10.5237 | -34.5339 | -20.234
10.5298 | -34.5344 | -20.2441
10.536 | -34.535 | -20.2548
10.5421 | -34.5355 | -20.2647
10.5482 | -34.536 | -20.2754
10.5543 | -34.5366 | -20.2858
10.5604 | -34.5371 | -20.2958
10.5665 | -34.5377 | -20.306
10.5726 | -34.5381 | -20.3162
10.5787 | -34.5387 | -20.3267
10.5848 | -34.5393 | -20.3367
10.5909 | -34.5398 | -20.3471
10.597 | -34.5403 | -20.3575
10.6031 | -34.5408 | -20.368
10.6092 | -34.5414 | -20.3786
10.6153 | -34.5419 | -20.3893
10.6214 | -34.5424 | -20.3995
10.6275 | -34.543 | -20.4097
10.6336 | -34.5435 | -20.4205
10.6397 | -34.544 | -20.4307
10.6458 | -34.5445 | -20.4408
10.6519 | -34.545 | -20.4511
10.658 | -34.5455 | -20.4611
10.6641 | -34.5461 | -20.4715
10.6702 | -34.5466 | -20.4821
10.6764 | -34.5472 | -20.4923
10.6825 | -34.5478 | -20.5028
10.6886 | -34.5483 | -20.5133
10.6947 | -34.5488 | -20.5234
10.7008 | -34.5494 | -20.5336
10.7069 | -34.5499 | -20.544
10.713 | -34.5505 | -20.5543
10.7191 | -34.551 | -20.5647
10.7252 | -34.5516 | -20.5747
10.7313 | -34.5521 | -20.5851
10.7374 | -34.5527 | -20.5952
10.7435 | -34.5532 | -20.6057
10.7496 | -34.5537 | -20.6159
10.7557 | -34.5542 | -20.6262
10.7618 | -34.5548 | -20.6364
10.7679 | -34.5553 | -20.6467
10.774 | -34.5558 | -20.6569
10.7801 | -34.5563 | -20.6677
10.7862 | -34.5569 | -20.6779
10.7923 | -34.5575 | -20.6882
10.7984 | -34.558 | -20.6984
10.8045 | -34.5585 | -20.7088
10.8106 | -34.5591 | -20.7191
10.8168 | -34.5596 | -20.7295
10.8229 | -34.5601 | -20.7398
10.829 | -34.5607 | -20.7502
10.8351 | -34.5612 | -20.7604
10.8412 | -34.5617 | -20.7706
10.8473 | -34.5622 | -20.7806
10.8534 | -34.5627 | -20.7911
10.8595 | -34.5633 | -20.8012
10.8656 | -34.5638 | -20.8112
10.8717 | -34.5644 | -20.8211
10.8778 | -34.5649 | -20.8313
10.8839 | -34.5654 | -20.8414
10.89 | -34.566 | -20.8514
10.8961 | -34.5665 | -20.8617
10.9022 | -34.5671 | -20.8719
10.9083 | -34.5676 | -20.8824
10.9144 | -34.5682 | -20.8925
10.9205 | -34.5688 | -20.9027
10.9266 | -34.5694 | -20.9133
10.9327 | -34.5699 | -20.9235
10.9388 | -34.5704 | -20.9337
10.9449 | -34.571 | -20.9439
10.951 | -34.5715 | -20.9541
10.9571 | -34.572 | -20.9644
10.9633 | -34.5726 | -20.975
10.9694 | -34.5731 | -20.9855
10.9755 | -34.5737 | -20.9957
10.9816 | -34.5742 | -21.006
10.9877 | -34.5748 | -21.0162
10.9938 | -34.5753 | -21.0267
10.9999 | -34.5759 | -21.0367
11.006 | -34.5764 | -21.0468
11.0121 | -34.5769 | -21.0574
11.0182 | -34.5775 | -21.0672
11.0243 | -34.578 | -21.0774
11.0304 | -34.5785 | -21.0875
11.0365 | -34.5791 | -21.0976
11.0426 | -34.5796 | -21.1076
11.0487 | -34.5802 | -21.1179
11.0548 | -34.5808 | -21.1281
11.0609 | -34.5813 | -21.1386
11.067 | -34.5818 | -21.149
11.0731 | -34.5824 | -21.1592
11.0792 | -34.5829 | -21.1693
11.0853 | -34.5835 | -21.1792
11.0914 | -34.584 | -21.1893
11.0975 | -34.5846 | -21.1997
11.1037 | -34.5851 | -21.2096
11.1098 | -34.5856 | -21.22
11.1159 | -34.5862 | -21.2303
11.122 | -34.5868 | -21.2408
11.1281 | -34.5873 | -21.2507
11.1342 | -34.5878 | -21.2606
11.1403 | -34.5883 | -21.271
11.1464 | -34.5889 | -21.2811
11.1525 | -34.5894 | -21.2915
11.1586 | -34.59 | -21.3023
11.1647 | -34.5906 | -21.3122
11.1708 | -34.5911 | -21.3223
11.1769 | -34.5916 | -21.3319
11.183 | -34.5922 | -21.342
11.1891 | -34.5927 | -21.3524
11.1952 | -34.5932 | -21.363
11.2013 | -34.5938 | -21.3731
11.2074 | -34.5943 | -21.383
11.2135 | -34.5948 | -21.3932
11.2196 | -34.5953 | -21.4032
11.2257 | -34.5958 | -21.4132
11.2318 | -34.5963 | -21.4237
11.2379 | -34.5967 | -21.4337
11.244 | -34.5973 | -21.4437
11.2502 | -34.5978 | -21.4538
11.2563 | -34.5984 | -21.464
11.2624 | -34.5989 | -21.4742
11.2685 | -34.5994 | -21.4842
11.2746 | -34.5999 | -21.494
11.2807 | -34.6004 | -21.5046
11.2868 | -34.6009 | -21.5151
11.2929 | -34.6016 | -21.5249
11.299 | -34.6021 | -21.5353
11.3051 | -34.6025 | -21.5453
11.3112 | -34.6031 | -21.5556
11.3173 | -34.6037 | -21.5658
11.3234 | -34.6042 | -21.576
11.3295 | -34.6047 | -21.5859
11.3356 | -34.6052 | -21.5959
11.3417 | -34.6058 | -21.6062
11.3478 | -34.6063 | -21.6163
11.3539 | -34.6068 | -21.6263
11.36 | -34.6073 | -21.6363
11.3661 | -34.6079 | -21.6463
11.3722 | -34.6084 | -21.6563
11.3783 | -34.609 | -21.6663
11.3844 | -34.6094 | -21.6765
11.3906 | -34.61 | -21.6869
11.3967 | -34.6105 | -21.6968
11.4028 | -34.611 | -21.7071
11.4089 | -34.6115 | -21.717
11.415 | -34.612 | -21.727
11.4211 | -34.6126 | -21.7369
11.4272 | -34.6131 | -21.7471
11.4333 | -34.6136 | -21.7572
11.4394 | -34.6141 | -21.7676
11.4455 | -34.6146 | -21.7777
11.4516 | -34.6152 | -21.7877
11.4577 | -34.6157 | -21.798
11.4638 | -34.6163 | -21.8076
11.4699 | -34.6168 | -21.818
11.476 | -34.6173 | -21.8286
11.4821 | -34.6179 | -21.8383
11.4882 | -34.6184 | -21.8488
11.4943 | -34.6189 | -21.8587
11.5004 | -34.6195 | -21.8687
11.5065 | -34.6201 | -21.8788
11.5126 | -34.6206 | -21.8886
11.5187 | -34.6212 | -21.8986
11.5248 | -34.6217 | -21.9087
11.5309 | -34.6222 | -21.9188
11.5371 | -34.6227 | -21.9287
11.5432 | -34.6232 | -21.939
11.5493 | -34.6237 | -21.9492
11.5554 | -34.6243 | -21.9591
11.5615 | -34.6248 | -21.9694
11.5676 | -34.6254 | -21.9797
11.5737 | -34.6259 | -21.9894
11.5798 | -34.6264 | -21.9992
11.5859 | -34.627 | -22.0091
11.592 | -34.6275 | -22.0192
11.5981 | -34.6281 | -22.0291
11.6042 | -34.6286 | -22.0392
11.6103 | -34.6291 | -22.0491
11.6164 | -34.6297 | -22.059
11.6225 | -34.6302 | -22.069
11.6286 | -34.6307 | -22.0792
11.6347 | -34.6313 | -22.0895
11.6408 | -34.6318 | -22.0995
11.6469 | -34.6324 | -22.1098
11.653 | -34.6329 | -22.1196
11.6591 | -34.6335 | -22.1297
11.6652 | -34.634 | -22.1398
11.6713 | -34.6345 | -22.1498
11.6775 | -34.6351 | -22.1596
11.6836 | -34.6356 | -22.1698
11.6897 | -34.6362 | -22.1796
11.6958 | -34.6367 | -22.19
11.7019 | -34.6373 | -22.2001
11.708 | -34.6378 | -22.2103
11.7141 | -34.6384 | -22.2204
11.7202 | -34.6389 | -22.2301
11.7263 | -34.6394 | -22.2403
11.7324 | -34.64 | -22.2504
11.7385 | -34.6405 | -22.2606
11.7446 | -34.641 | -22.2708
11.7507 | -34.6416 | -22.2808
11.7568 | -34.6421 | -22.2909
11.7629 | -34.6427 | -22.3009
11.769 | -34.6432 | -22.3109
11.7751 | -34.6438 | -22.3206
11.7812 | -34.6443 | -22.3304
11.7873 | -34.6449 | -22.3407
11.7934 | -34.6455 | -22.3507
11.7995 | -34.646 | -22.3604
11.8056 | -34.6465 | -22.3708
11.8117 | -34.6469 | -22.3808
11.8178 | -34.6475 | -22.3908
11.824 | -34.6481 | -22.4006
11.8301 | -34.6487 | -22.4102
11.8362 | -34.6492 | -22.4199
11.8423 | -34.6497 | -22.4299
11.8484 | -34.6503 | -22.4399
11.8545 | -34.6508 | -22.4501
11.8606 | -34.6514 | -22.4601
11.8667 | -34.652 | -22.4702
11.8728 | -34.6525 | -22.4801
11.8789 | -34.653 | -22.4905
11.885 | -34.6535 | -22.5004
11.8911 | -34.6541 | -22.5106
11.8972 | -34.6546 | -22.5202
11.9033 | -34.6551 | -22.5297
11.9094 | -34.6556 | -22.5393
11.9155 | -34.6561 | -22.5493
11.9216 | -34.6567 | -22.5591
11.9277 | -34.6573 | -22.5688
11.9338 | -34.6578 | -22.5789
11.9399 | -34.6583 | -22.5891
11.946 | -34.6589 | -22.5991
11.9521 | -34.6595 | -22.609
11.9582 | -34.66 | -22.6187
11.9644 | -34.6606 | -22.6285
11.9705 | -34.6611 | -22.6386
11.9766 | -34.6617 | -22.6485
11.9827 | -34.6622 | -22.6583
11.9888 | -34.6627 | -22.668
11.9949 | -34.6632 | -22.6779
12.001 | -34.6638 | -22.6878
12.0071 | -34.6644 | -22.6975
12.0132 | -34.6649 | -22.7071
12.0193 | -34.6655 | -22.7173
12.0254 | -34.666 | -22.727
12.0315 | -34.6666 | -22.7367
12.0376 | -34.6671 | -22.7464
12.0437 | -34.6676 | -22.756
12.0498 | -34.6682 | -22.7662
12.0559 | -34.6687 | -22.7762
12.062 | -34.6693 | -22.7859
12.0681 | -34.6698 | -22.7958
12.0742 | -34.6704 | -22.8059
12.0803 | -34.6709 | -22.8156
12.0864 | -34.6715 | -22.8255
12.0925 | -34.6721 | -22.8354
12.0986 | -34.6726 | -22.8449
12.1047 | -34.6732 | -22.8547
12.1109 | -34.6738 | -22.8648
12.117 | -34.6743 | -22.8745
12.1231 | -34.6749 | -22.8844
12.1292 | -34.6754 | -22.8944
12.1353 | -34.676 | -22.9039
12.1414 | -34.6766 | -22.914
12.1475 | -34.6771 | -22.9239
12.1536 | -34.6776 | -22.9338
12.1597 | -34.6782 | -22.9437
12.1658 | -34.6787 | -22.9536
12.1719 | -34.6793 | -22.9635
12.178 | -34.6799 | -22.9735
12.1841 | -34.6804 | -22.9835
12.1902 | -34.681 | -22.9933
12.1963 | -34.6816 | -23.0031
12.2024 | -34.6822 | -23.0128
12.2085 | -34.6827 | -23.0227
12.2146 | -34.6832 | -23.0323
12.2207 | -34.6838 | -23.0421
12.2268 | -34.6843 | -23.0519
12.2329 | -34.6849 | -23.0617
12.239 | -34.6854 | -23.0712
12.2451 | -34.686 | -23.0808
12.2513 | -34.6866 | -23.0907
12.2574 | -34.6871 | -23.1008
12.2635 | -34.6877 | -23.1102
12.2696 | -34.6882 | -23.1199
12.2757 | -34.6888 | -23.1293
12.2818 | -34.6893 | -23.1389
12.2879 | -34.6898 | -23.1485
12.294 | -34.6904 | -23.1581
12.3001 | -34.6909 | -23.168
12.3062 | -34.6915 | -23.1778
12.3123 | -34.6921 | -23.1876
12.3184 | -34.6926 | -23.1971
12.3245 | -34.6932 | -23.207
12.3306 | -34.6938 | -23.2165
12.3367 | -34.6943 | -23.2261
12.3428 | -34.6949 | -23.2356
12.3489 | -34.6954 | -23.2449
12.355 | -34.696 | -23.2545
12.3611 | -34.6965 | -23.2642
12.3672 | -34.6971 | -23.2738
12.3733 | -34.6976 | -23.2839
12.3794 | -34.6982 | -23.2931
12.3855 | -34.6987 | -23.3026
12.3916 | -34.6992 | -23.3126
12.3978 | -34.6998 | -23.3226
12.4039 | -34.7003 | -23.3323
12.41 | -34.7009 | -23.3422
12.4161 | -34.7015 | -23.3515
12.4222 | -34.7021 | -23.3611
12.4283 | -34.7026 | -23.3708
12.4344 | -34.7032 | -23.3805
12.4405 | -34.7037 | -23.3902
12.4466 | -34.7043 | -23.3997
12.4527 | -34.7048 | -23.4094
12.4588 | -34.7054 | -23.4191
12.4649 | -34.706 | -23.429
12.471 | -34.7065 | -23.4392
12.4771 | -34.7071 | -23.4482
12.4832 | -34.7076 | -23.4578
12.4893 | -34.7082 | -23.4672
12.4954 | -34.7088 | -23.4767
12.5015 | -34.7094 | -23.4861
12.5076 | -34.7099 | -23.4959
12.5137 | -34.7104 | -23.506
12.5198 | -34.711 | -23.5153
12.5259 | -34.7116 | -23.5248
12.532 | -34.7122 | -23.5347
12.5382 | -34.7127 | -23.5441
12.5443 | -34.7133 | -23.5539
12.5504 | -34.7138 | -23.5636
12.5565 | -34.7145 | -23.5729
12.5626 | -34.715 | -23.5829
12.5687 | -34.7156 | -23.5927
12.5748 | -34.7161 | -23.6023
12.5809 | -34.7167 | -23.6115
12.587 | -34.7172 | -23.6215
12.5931 | -34.7178 | -23.6311
12.5992 | -34.7183 | -23.641
12.6053 | -34.7189 | -23.6505
12.6114 | -34.7195 | -23.6601
12.6175 | -34.7201 | -23.6699
12.6236 | -34.7206 | -23.6792
12.6297 | -34.7212 | -23.6886
12.6358 | -34.7218 | -23.6979
12.6419 | -34.7224 | -23.7078
12.648 | -34.7229 | -23.7177
12.6541 | -34.7235 | -23.7272
12.6602 | -34.7241 | -23.7363
12.6663 | -34.7246 | -23.7458
12.6724 | -34.7252 | -23.7554
12.6785 | -34.7257 | -23.7653
12.6847 | -34.7263 | -23.7749
12.6908 | -34.7269 | -23.7844
12.6969 | -34.7275 | -23.7939
12.703 | -34.7281 | -23.8036
12.7091 | -34.7286 | -23.8133
12.7152 | -34.7292 | -23.8228
12.7213 | -34.7297 | -23.8325
12.7274 | -34.7303 | -23.8421
12.7335 | -34.7309 | -23.851
12.7396 | -34.7314 | -23.8602
12.7457 | -34.732 | -23.8698
12.7518 | -34.7325 | -23.8794
12.7579 | -34.7331 | -23.8893
12.764 | -34.7336 | -23.8989
12.7701 | -34.7341 | -23.9081
12.7762 | -34.7347 | -23.9177
12.7823 | -34.7352 | -23.9273
12.7884 | -34.7358 | -23.9368
12.7945 | -34.7363 | -23.946
12.8006 | -34.7369 | -23.9548
12.8067 | -34.7375 | -23.9647
12.8128 | -34.738 | -23.9739
12.8189 | -34.7385 | -23.9831
12.8251 | -34.739 | -23.9924
12.8312 | -34.7395 | -24.002
12.8373 | -34.7401 | -24.0114
12.8434 | -34.7407 | -24.0208
12.8495 | -34.7412 | -24.0302
12.8556 | -34.7418 | -24.0396
12.8617 | -34.7423 | -24.0488
12.8678 | -34.7428 | -24.0582
12.8739 | -34.7434 | -24.0673
12.88 | -34.7439 | -24.0769
12.8861 | -34.7445 | -24.0865
12.8922 | -34.745 | -24.0959
12.8983 | -34.7456 | -24.105
12.9044 | -34.7462 | -24.1144
12.9105 | -34.7468 | -24.1237
12.9166 | -34.7474 | -24.1335
12.9227 | -34.7479 | -24.1435
12.9288 | -34.7485 | -24.1531
12.9349 | -34.7491 | -24.1629
12.941 | -34.7497 | -24.1724
12.9471 | -34.7503 | -24.1819
12.9532 | -34.7509 | -24.1913
12.9593 | -34.7515 | -24.201
12.9654 | -34.752 | -24.2106
12.9716 | -34.7526 | -24.2201
12.9777 | -34.7531 | -24.2291
12.9838 | -34.7536 | -24.2384
12.9899 | -34.7542 | -24.2479
12.996 | -34.7547 | -24.2576
13.0021 | -34.7552 | -24.267
13.0082 | -34.7557 | -24.2765
13.0143 | -34.7563 | -24.2855
13.0204 | -34.7568 | -24.2946
13.0265 | -34.7574 | -24.3037
13.0326 | -34.7579 | -24.3134
13.0387 | -34.7585 | -24.3231
13.0448 | -34.759 | -24.3324
13.0509 | -34.7596 | -24.3418
13.057 | -34.7602 | -24.3508
13.0631 | -34.7609 | -24.3601
13.0692 | -34.7614 | -24.3694
13.0753 | -34.762 | -24.3787
13.0814 | -34.7625 | -24.3879
13.0875 | -34.7631 | -24.3968
13.0936 | -34.7637 | -24.406
13.0997 | -34.7642 | -24.4151
13.1058 | -34.7648 | -24.4239
13.112 | -34.7653 | -24.4333
13.1181 | -34.7659 | -24.4426
13.1242 | -34.7664 | -24.452
13.1303 | -34.767 | -24.4616
13.1364 | -34.7675 | -24.471
13.1425 | -34.768 | -24.4802
13.1486 | -34.7686 | -24.4894
13.1547 | -34.7692 | -24.4984
13.1608 | -34.7697 | -24.5077
13.1669 | -34.7703 | -24.517
13.173 | -34.7708 | -24.5262
13.1791 | -34.7714 | -24.5356
13.1852 | -34.7719 | -24.5453
13.1913 | -34.7725 | -24.5546
13.1974 | -34.773 | -24.5639
13.2035 | -34.7735 | -24.5732
13.2096 | -34.7741 | -24.5825
13.2157 | -34.7747 | -24.5918
13.2218 | -34.7753 | -24.6012
13.2279 | -34.7759 | -24.6104
13.234 | -34.7764 | -24.6198
13.2401 | -34.777 | -24.629
13.2462 | -34.7775 | -24.6382
13.2524 | -34.7781 | -24.6474
13.2585 | -34.7786 | -24.6566
13.2646 | -34.7792 | -24.6661
13.2707 | -34.7797 | -24.6754
13.2768 | -34.7803 | -24.6851
13.2829 | -34.7808 | -24.6945
13.289 | -34.7814 | -24.7037
13.2951 | -34.782 | -24.7128
13.3012 | -34.7825 | -24.7221
13.3073 | -34.783 | -24.7317
13.3134 | -34.7835 | -24.7408
13.3195 | -34.7841 | -24.7499
13.3256 | -34.7847 | -24.7593
13.3317 | -34.7852 | -24.7684
13.3378 | -34.7858 | -24.7778
13.3439 | -34.7863 | -24.7871
13.35 | -34.7869 | -24.7964
13.3561 | -34.7874 | -24.8058
13.3622 | -34.7879 | -24.8151
13.3683 | -34.7885 | -24.824
13.3744 | -34.789 | -24.8332
13.3805 | -34.7896 | -24.8426
13.3866 | -34.7901 | -24.8517
13.3927 | -34.7906 | -24.8608
13.3989 | -34.7912 | -24.8698
13.405 | -34.7918 | -24.8789
13.4111 | -34.7924 | -24.8882
13.4172 | -34.7929 | -24.8971
13.4233 | -34.7934 | -24.9066
13.4294 | -34.794 | -24.9156
13.4355 | -34.7945 | -24.9255
13.4416 | -34.795 | -24.9348
13.4477 | -34.7955 | -24.9441
13.4538 | -34.7961 | -24.9533
13.4599 | -34.7967 | -24.9627
13.466 | -34.7972 | -24.9722
13.4721 | -34.7978 | -24.9815
13.4782 | -34.7984 | -24.9907
13.4843 | -34.799 | -25
13.4904 | -34.7995 | -25.0087
13.4965 | -34.8 | -25.0184
13.5026 | -34.8006 | -25.0275
13.5087 | -34.8012 | -25.0364
13.5148 | -34.8017 | -25.0458
13.5209 | -34.8022 | -25.0548
13.527 | -34.8028 | -25.0644
13.5331 | -34.8033 | -25.0736
13.5393 | -34.8039 | -25.0828
13.5454 | -34.8043 | -25.0922
13.5515 | -34.8049 | -25.1013
13.5576 | -34.8055 | -25.1108
13.5637 | -34.806 | -25.1199
13.5698 | -34.8066 | -25.1294
13.5759 | -34.8071 | -25.1381
13.582 | -34.8076 | -25.1475
13.5881 | -34.8081 | -25.1567
13.5942 | -34.8087 | -25.1659
13.6003 | -34.8093 | -25.1753
13.6064 | -34.8098 | -25.1846
13.6125 | -34.8104 | -25.1941
13.6186 | -34.811 | -25.2033
13.6247 | -34.8115 | -25.2124
13.6308 | -34.812 | -25.2214
13.6369 | -34.8126 | -25.2306
13.643 | -34.8132 | -25.2396
13.6491 | -34.8137 | -25.2489
13.6552 | -34.8143 | -25.258
13.6613 | -34.8148 | -25.2673
13.6674 | -34.8154 | -25.2766
13.6735 | -34.8159 | -25.286
13.6796 | -34.8165 | -25.295
13.6858 | -34.8169 | -25.3044
13.6919 | -34.8175 | -25.3137
13.698 | -34.8181 | -25.323
13.7041 | -34.8186 | -25.3325
13.7102 | -34.8192 | -25.3418
13.7163 | -34.8198 | -25.3512
13.7224 | -34.8203 | -25.3604
13.7285 | -34.8209 | -25.3697
13.7346 | -34.8214 | -25.3789
13.7407 | -34.822 | -25.3882
13.7468 | -34.8225 | -25.3974
13.7529 | -34.823 | -25.4065
13.759 | -34.8235 | -25.4158
13.7651 | -34.8241 | -25.4252
13.7712 | -34.8246 | -25.4346
13.7773 | -34.8252 | -25.4439
13.7834 | -34.8257 | -25.4531
13.7895 | -34.8264 | -25.4621
13.7956 | -34.8269 | -25.4715
13.8017 | -34.8275 | -25.4806
13.8078 | -34.828 | -25.4899
13.8139 | -34.8286 | -25.4989
13.82 | -34.8292 | -25.5078
13.8262 | -34.8297 | -25.5174
13.8323 | -34.8303 | -25.5268
13.8384 | -34.8308 | -25.5362
13.8445 | -34.8314 | -25.5456
13.8506 | -34.832 | -25.555
13.8567 | -34.8325 | -25.5644
13.8628 | -34.833 | -25.5737
13.8689 | -34.8335 | -25.583
13.875 | -34.8341 | -25.5922
13.8811 | -34.8347 | -25.6014
13.8872 | -34.8353 | -25.6106
13.8933 | -34.8358 | -25.6198
13.8994 | -34.8364 | -25.6287
13.9055 | -34.837 | -25.6379
13.9116 | -34.8376 | -25.6468
13.9177 | -34.8381 | -25.6557
13.9238 | -34.8387 | -25.6651
13.9299 | -34.8393 | -25.6744
13.936 | -34.8398 | -25.6834
13.9421 | -34.8404 | -25.6926
13.9482 | -34.841 | -25.7019
13.9543 | -34.8415 | -25.7109
13.9604 | -34.8421 | -25.7202
13.9665 | -34.8427 | -25.7295
13.9727 | -34.8432 | -25.7385
13.9788 | -34.8437 | -25.7479
13.9849 | -34.8442 | -25.7572
13.991 | -34.8447 | -25.7667
13.9971 | -34.8453 | -25.7762
14.0032 | -34.8458 | -25.7856
14.0093 | -34.8463 | -25.7946
14.0154 | -34.8468 | -25.8041
14.0215 | -34.8474 | -25.813
14.0276 | -34.8479 | -25.822
14.0337 | -34.8485 | -25.8315
14.0398 | -34.8491 | -25.8404
14.0459 | -34.8496 | -25.8499
14.052 | -34.8501 | -25.8591
14.0581 | -34.8506 | -25.868
14.0642 | -34.8512 | -25.8774
14.0703 | -34.8517 | -25.8866
14.0764 | -34.8523 | -25.8962
14.0825 | -34.8529 | -25.9054
14.0886 | -34.8534 | -25.9141
14.0947 | -34.8539 | -25.9231
14.1008 | -34.8545 | -25.9326
14.1069 | -34.855 | -25.9417
14.1131 | -34.8555 | -25.9512
14.1192 | -34.8561 | -25.96
14.1253 | -34.8566 | -25.969
14.1314 | -34.8572 | -25.9781
14.1375 | -34.8577 | -25.9875
14.1436 | -34.8583 | -25.9966
14.1497 | -34.8589 | -26.0055
14.1558 | -34.8595 | -26.0149
14.1619 | -34.86 | -26.024
14.168 | -34.8606 | -26.033
14.1741 | -34.8612 | -26.042
14.1802 | -34.8617 | -26.051
14.1863 | -34.8623 | -26.0603
14.1924 | -34.8628 | -26.0695
14.1985 | -34.8634 | -26.0785
14.2046 | -34.8639 | -26.0874
14.2107 | -34.8645 | -26.0961
14.2168 | -34.8651 | -26.1057
14.2229 | -34.8656 | -26.115
14.229 | -34.8662 | -26.1241
14.2351 | -34.8667 | -26.133
14.2412 | -34.8673 | -26.142
14.2473 | -34.8679 | -26.1511
14.2534 | -34.8685 | -26.1602
14.2596 | -34.869 | -26.1693
14.2657 | -34.8695 | -26.1787
14.2718 | -34.8701 | -26.188
14.2779 | -34.8707 | -26.1969
14.284 | -34.8712 | -26.2059
14.2901 | -34.8718 | -26.2148
14.2962 | -34.8724 | -26.2237
14.3023 | -34.8729 | -26.233
14.3084 | -34.8735 | -26.2424
14.3145 | -34.8741 | -26.2515
14.3206 | -34.8746 | -26.2607
14.3267 | -34.8751 | -26.2698
14.3328 | -34.8757 | -26.2788
14.3389 | -34.8762 | -26.2877
14.345 | -34.8767 | -26.297
14.3511 | -34.8773 | -26.3058
14.3572 | -34.8778 | -26.3152
14.3633 | -34.8784 | -26.3242
14.3694 | -34.8789 | -26.3333
14.3755 | -34.8794 | -26.3421
14.3816 | -34.88 | -26.3511
14.3877 | -34.8806 | -26.36
14.3938 | -34.8812 | -26.3693
14.4 | -34.8817 | -26.3783
14.4061 | -34.8822 | -26.3874
14.4122 | -34.8828 | -26.3967
14.4183 | -34.8833 | -26.4058
14.4244 | -34.8838 | -26.4148
14.4305 | -34.8844 | -26.4242
14.4366 | -34.885 | -26.4334
14.4427 | -34.8856 | -26.4427
14.4488 | -34.8861 | -26.4514
14.4549 | -34.8866 | -26.4604
14.461 | -34.8872 | -26.4696
14.4671 | -34.8878 | -26.4787
14.4732 | -34.8883 | -26.4874
14.4793 | -34.8889 | -26.4972
14.4854 | -34.8895 | -26.5059
14.4915 | -34.8901 | -26.5147
14.4976 | -34.8907 | -26.5239
14.5037 | -34.8913 | -26.5331
14.5098 | -34.8918 | -26.5419
14.5159 | -34.8924 | -26.5512
14.522 | -34.8929 | -26.5604
14.5281 | -34.8935 | -26.5692
14.5342 | -34.894 | -26.5784
14.5403 | -34.8946 | -26.5876
14.5465 | -34.8952 | -26.5971
14.5526 | -34.8958 | -26.6064
14.5587 | -34.8963 | -26.6155
14.5648 | -34.8969 | -26.6243
14.5709 | -34.8974 | -26.6336
14.577 | -34.898 | -26.6427
14.5831 | -34.8985 | -26.6519
14.5892 | -34.8992 | -26.6605
14.5953 | -34.8997 | -26.6692
14.6014 | -34.9003 | -26.6782
14.6075 | -34.9009 | -26.687
14.6136 | -34.9014 | -26.6958
14.6197 | -34.902 | -26.7049
14.6258 | -34.9025 | -26.7141
14.6319 | -34.9031 | -26.7228
14.638 | -34.9037 | -26.7319
14.6441 | -34.9043 | -26.7404
14.6502 | -34.9048 | -26.7492
14.6563 | -34.9054 | -26.7581
14.6624 | -34.906 | -26.767
14.6685 | -34.9065 | -26.7755
14.6746 | -34.9071 | -26.7846
14.6807 | -34.9076 | -26.7936
14.6869 | -34.9082 | -26.8025
14.693 | -34.9088 | -26.8113
14.6991 | -34.9093 | -26.8201
14.7052 | -34.9099 | -26.8289
14.7113 | -34.9104 | -26.8382
14.7174 | -34.911 | -26.847
14.7235 | -34.9116 | -26.8559
14.7296 | -34.9121 | -26.8648
14.7357 | -34.9127 | -26.8735
14.7418 | -34.9133 | -26.8821
14.7479 | -34.9139 | -26.8911
14.754 | -34.9144 | -26.9
14.7601 | -34.915 | -26.9091
14.7662 | -34.9156 | -26.9181
14.7723 | -34.9161 | -26.9268
14.7784 | -34.9167 | -26.9358
14.7845 | -34.9173 | -26.9448
14.7906 | -34.9179 | -26.9538
14.7967 | -34.9184 | -26.9627
14.8028 | -34.9189 | -26.9715
14.8089 | -34.9195 | -26.9802
14.815 | -34.9201 | -26.9889
14.8211 | -34.9206 | -26.9977
14.8272 | -34.9211 | -27.0067
14.8334 | -34.9217 | -27.0155
14.8395 | -34.9223 | -27.0244
14.8456 | -34.9229 | -27.0332
14.8517 | -34.9235 | -27.0423
14.8578 | -34.924 | -27.0511
14.8639 | -34.9246 | -27.0602
14.87 | -34.9252 | -27.0691
14.8761 | -34.9258 | -27.0777
14.8822 | -34.9263 | -27.0866
14.8883 | -34.9269 | -27.0953
14.8944 | -34.9275 | -27.1043
14.9005 | -34.9281 | -27.1129
14.9066 | -34.9286 | -27.1219
14.9127 | -34.9292 | -27.1305
14.9188 | -34.9298 | -27.1396
14.9249 | -34.9303 | -27.1485
14.931 | -34.9309 | -27.1571
14.9371 | -34.9315 | -27.1659
14.9432 | -34.9321 | -27.1748
14.9493 | -34.9326 | -27.1836
14.9554 | -34.9332 | -27.1924
14.9615 | -34.9338 | -27.2013
14.9676 | -34.9344 | -27.2104
14.9738 | -34.9349 | -27.2191
14.9799 | -34.9355 | -27.2283
14.986 | -34.9361 | -27.2366
14.9921 | -34.9367 | -27.2453
14.9982 | -34.9372 | -27.2545
15.0043 | -34.9378 | -27.2633
15.0104 | -34.9383 | -27.2722
15.0165 | -34.939 | -27.2812
15.0226 | -34.9395 | -27.2896
15.0287 | -34.9401 | -27.2989
15.0348 | -34.9407 | -27.3078
15.0409 | -34.9412 | -27.317
15.047 | -34.9418 | -27.3261
15.0531 | -34.9425 | -27.335
15.0592 | -34.9431 | -27.344
15.0653 | -34.9437 | -27.3527
15.0714 | -34.9443 | -27.3613
15.0775 | -34.9449 | -27.3698
15.0836 | -34.9454 | -27.3787
15.0897 | -34.9461 | -27.3874
15.0958 | -34.9467 | -27.396
15.1019 | -34.9472 | -27.4048
15.108 | -34.9478 | -27.4136
15.1141 | -34.9483 | -27.4225
15.1203 | -34.9489 | -27.4313
15.1264 | -34.9496 | -27.4402
15.1325 | -34.9502 | -27.4489
15.1386 | -34.9508 | -27.4578
15.1447 | -34.9514 | -27.4667
15.1508 | -34.952 | -27.4758
15.1569 | -34.9526 | -27.4848
15.163 | -34.9533 | -27.4931
15.1691 | -34.9539 | -27.5018
15.1752 | -34.9545 | -27.5106
15.1813 | -34.955 | -27.5194
15.1874 | -34.9557 | -27.5284
15.1935 | -34.9563 | -27.5368
15.1996 | -34.9569 | -27.5458
15.2057 | -34.9575 | -27.5544
15.2118 | -34.9581 | -27.5631
15.2179 | -34.9587 | -27.5716
15.224 | -34.9592 | -27.5805
15.2301 | -34.9598 | -27.5889
15.2362 | -34.9605 | -27.5976
15.2423 | -34.961 | -27.6067
15.2484 | -34.9616 | -27.6153
15.2545 | -34.9622 | -27.6239
15.2607 | -34.9628 | -27.6328
15.2668 | -34.9635 | -27.6412
15.2729 | -34.9641 | -27.6498
15.279 | -34.9647 | -27.6586
15.2851 | -34.9653 | -27.6674
15.2912 | -34.9659 | -27.676
15.2973 | -34.9664 | -27.6847
15.3034 | -34.9671 | -27.6931
15.3095 | -34.9677 | -27.7016
15.3156 | -34.9682 | -27.7106
15.3217 | -34.9688 | -27.7189
15.3278 | -34.9695 | -27.7278
15.3339 | -34.9701 | -27.7361
15.34 | -34.9707 | -27.7446
15.3461 | -34.9713 | -27.7534
15.3522 | -34.9718 | -27.762
15.3583 | -34.9725 | -27.7706
15.3644 | -34.973 | -27.7794
15.3705 | -34.9737 | -27.7878
15.3766 | -34.9743 | -27.7968
15.3827 | -34.9748 | -27.8052
15.3888 | -34.9754 | -27.8136
15.3949 | -34.976 | -27.8221
15.401 | -34.9766 | -27.8309
15.4072 | -34.9773 | -27.8391
15.4133 | -34.9779 | -27.8478
15.4194 | -34.9785 | -27.8565
15.4255 | -34.9791 | -27.8651
15.4316 | -34.9797 | -27.8738
15.4377 | -34.9803 | -27.8823
15.4438 | -34.9809 | -27.8905
15.4499 | -34.9815 | -27.8991
15.456 | -34.982 | -27.9077
15.4621 | -34.9826 | -27.9162
15.4682 | -34.9833 | -27.925
15.4743 | -34.9839 | -27.9334
15.4804 | -34.9845 | -27.9421
15.4865 | -34.9851 | -27.9507
15.4926 | -34.9857 | -27.9594
15.4987 | -34.9863 | -27.9681
15.5048 | -34.9869 | -27.9767
15.5109 | -34.9875 | -27.9849
15.517 | -34.9882 | -27.9936
15.5231 | -34.9888 | -28.0022
15.5292 | -34.9894 | -28.0106
15.5353 | -34.99 | -28.0191
15.5414 | -34.9906 | -28.028
15.5476 | -34.9912 | -28.0367
15.5537 | -34.9918 | -28.0453
15.5598 | -34.9925 | -28.0537
15.5659 | -34.9931 | -28.0625
15.572 | -34.9937 | -28.0706
15.5781 | -34.9943 | -28.0793
15.5842 | -34.9949 | -28.0875
15.5903 | -34.9955 | -28.0958
15.5964 | -34.9961 | -28.1042
15.6025 | -34.9967 | -28.1126
15.6086 | -34.9973 | -28.1212
15.6147 | -34.9979 | -28.1298
15.6208 | -34.9986 | -28.138
15.6269 | -34.9993 | -28.1464
15.633 | -34.9998 | -28.155
15.6391 | -35.0004 | -28.1637
15.6452 | -35.0011 | -28.1722
15.6513 | -35.0017 | -28.1809
15.6574 | -35.0023 | -28.1891
15.6635 | -35.0029 | -28.1977
15.6696 | -35.0035 | -28.2061
15.6757 | -35.0042 | -28.215
15.6818 | -35.0047 | -28.2239
15.688 | -35.0054 | -28.2323
15.6941 | -35.006 | -28.2407
15.7002 | -35.0066 | -28.2492
15.7063 | -35.0072 | -28.2578
15.7124 | -35.0078 | -28.2664
15.7185 | -35.0084 | -28.2749
15.7246 | -35.0091 | -28.2832
15.7307 | -35.0097 | -28.2919
15.7368 | -35.0103 | -28.3006
15.7429 | -35.0109 | -28.309
15.749 | -35.0115 | -28.3176
15.7551 | -35.0121 | -28.3261
15.7612 | -35.0127 | -28.334
15.7673 | -35.0133 | -28.3423
15.7734 | -35.0139 | -28.3506
15.7795 | -35.0145 | -28.3595
15.7856 | -35.0152 | -28.3678
15.7917 | -35.0158 | -28.376
15.7978 | -35.0164 | -28.3844
15.8039 | -35.017 | -28.3928
15.81 | -35.0176 | -28.4012
15.8161 | -35.0182 | -28.4098
15.8222 | -35.0188 | -28.4182
15.8283 | -35.0195 | -28.4265
15.8345 | -35.0201 | -28.435
15.8406 | -35.0207 | -28.4434
15.8467 | -35.0214 | -28.4519
15.8528 | -35.022 | -28.4602
15.8589 | -35.0226 | -28.4686
15.865 | -35.0232 | -28.4769
15.8711 | -35.0239 | -28.4856
15.8772 | -35.0245 | -28.4942
15.8833 | -35.0251 | -28.503
15.8894 | -35.0257 | -28.5114
15.8955 | -35.0264 | -28.52
15.9016 | -35.027 | -28.5283
15.9077 | -35.0277 | -28.5368
15.9138 | -35.0283 | -28.5452
15.9199 | -35.0289 | -28.5537
15.926 | -35.0295 | -28.5625
15.9321 | -35.0302 | -28.5708
15.9382 | -35.0309 | -28.5791
15.9443 | -35.0315 | -28.5878
15.9504 | -35.0322 | -28.5962
15.9565 | -35.0328 | -28.6046
15.9626 | -35.0335 | -28.6132
15.9687 | -35.0341 | -28.6214
15.9749 | -35.0348 | -28.63
15.981 | -35.0354 | -28.6384
15.9871 | -35.036 | -28.6468
15.9932 | -35.0366 | -28.6552
15.9993 | -35.0373 | -28.6635
16.0054 | -35.0379 | -28.6717
16.0115 | -35.0386 | -28.6799
16.0176 | -35.0392 | -28.6885
16.0237 | -35.0398 | -28.6974
16.0298 | -35.0404 | -28.7053
16.0359 | -35.0411 | -28.7136
16.042 | -35.0417 | -28.7219
16.0481 | -35.0423 | -28.7303
16.0542 | -35.043 | -28.739
16.0603 | -35.0436 | -28.7473
16.0664 | -35.0443 | -28.7557
16.0725 | -35.0449 | -28.7638
16.0786 | -35.0456 | -28.7724
16.0847 | -35.0462 | -28.7804
16.0908 | -35.0468 | -28.7888
16.0969 | -35.0475 | -28.797
16.103 | -35.048 | -28.8056
16.1091 | -35.0487 | -28.814
16.1152 | -35.0493 | -28.8225
16.1214 | -35.0499 | -28.8309
16.1275 | -35.0506 | -28.8391
16.1336 | -35.0512 | -28.8475
16.1397 | -35.0519 | -28.8559
16.1458 | -35.0525 | -28.8641
16.1519 | -35.0531 | -28.8727
16.158 | -35.0538 | -28.8813
16.1641 | -35.0544 | -28.8898
16.1702 | -35.055 | -28.8982
16.1763 | -35.0556 | -28.9066
16.1824 | -35.0563 | -28.9146
16.1885 | -35.0569 | -28.923
16.1946 | -35.0576 | -28.9313
16.2007 | -35.0582 | -28.9402
16.2068 | -35.0588 | -28.9484
16.2129 | -35.0594 | -28.9568
16.219 | -35.06 | -28.9652
16.2251 | -35.0607 | -28.9736
16.2312 | -35.0612 | -28.982
16.2373 | -35.0619 | -28.9904
16.2434 | -35.0625 | -28.9986
16.2495 | -35.0632 | -29.0067
16.2556 | -35.0638 | -29.0151
16.2618 | -35.0644 | -29.0231
16.2679 | -35.0651 | -29.0316
16.274 | -35.0657 | -29.0401
16.2801 | -35.0664 | -29.0482
16.2862 | -35.067 | -29.0563
16.2923 | -35.0676 | -29.0648
16.2984 | -35.0682 | -29.0733
16.3045 | -35.0689 | -29.0817
16.3106 | -35.0696 | -29.0897
16.3167 | -35.0702 | -29.0981
16.3228 | -35.0708 | -29.1065
16.3289 | -35.0714 | -29.1148
16.335 | -35.0721 | -29.1234
16.3411 | -35.0727 | -29.1316
16.3472 | -35.0733 | -29.14
16.3533 | -35.0739 | -29.1484
16.3594 | -35.0745 | -29.1565
16.3655 | -35.0752 | -29.1651
16.3716 | -35.0758 | -29.1733
16.3777 | -35.0764 | -29.1816
16.3838 | -35.077 | -29.1898
16.3899 | -35.0777 | -29.1983
16.396 | -35.0783 | -29.2068
16.4021 | -35.0789 | -29.2151
16.4083 | -35.0796 | -29.2234
16.4144 | -35.0803 | -29.2316
16.4205 | -35.0809 | -29.2401
16.4266 | -35.0815 | -29.2483
16.4327 | -35.0822 | -29.2564
16.4388 | -35.0828 | -29.2647
16.4449 | -35.0834 | -29.2731
16.451 | -35.0841 | -29.2816
16.4571 | -35.0847 | -29.29
16.4632 | -35.0853 | -29.2981
16.4693 | -35.0859 | -29.3065
16.4754 | -35.0866 | -29.3149
16.4815 | -35.0872 | -29.3234
16.4876 | -35.0878 | -29.3316
16.4937 | -35.0884 | -29.3399
16.4998 | -35.0891 | -29.3482
16.5059 | -35.0898 | -29.3565
16.512 | -35.0904 | -29.3652
16.5181 | -35.0911 | -29.3737
16.5242 | -35.0917 | -29.3821
16.5303 | -35.0923 | -29.3901
16.5364 | -35.093 | -29.3983
16.5425 | -35.0936 | -29.4067
16.5487 | -35.0942 | -29.4148
16.5548 | -35.0948 | -29.4231
16.5609 | -35.0955 | -29.4316
16.567 | -35.0961 | -29.4397
16.5731 | -35.0967 | -29.4481
16.5792 | -35.0973 | -29.4563
16.5853 | -35.098 | -29.4645
16.5914 | -35.0986 | -29.4729
16.5975 | -35.0993 | -29.4814
16.6036 | -35.0999 | -29.4897
16.6097 | -35.1006 | -29.498
16.6158 | -35.1012 | -29.5064
16.6219 | -35.1018 | -29.5146
16.628 | -35.1024 | -29.523
16.6341 | -35.103 | -29.5311
16.6402 | -35.1037 | -29.539
16.6463 | -35.1043 | -29.5473
16.6524 | -35.1049 | -29.5554
16.6585 | -35.1055 | -29.5637
16.6646 | -35.1062 | -29.5722
16.6707 | -35.1068 | -29.5803
16.6768 | -35.1075 | -29.5885
16.6829 | -35.1081 | -29.5967
16.689 | -35.1087 | -29.6051
16.6952 | -35.1094 | -29.6135
16.7013 | -35.11 | -29.6219
16.7074 | -35.1106 | -29.6301
16.7135 | -35.1113 | -29.6384
16.7196 | -35.1119 | -29.6468
16.7257 | -35.1125 | -29.6548
16.7318 | -35.1131 | -29.6631
16.7379 | -35.1137 | -29.6716
16.744 | -35.1144 | -29.6796
16.7501 | -35.115 | -29.688
16.7562 | -35.1157 | -29.6963
16.7623 | -35.1163 | -29.7047
16.7684 | -35.1169 | -29.7131
16.7745 | -35.1176 | -29.7211
16.7806 | -35.1182 | -29.7295
16.7867 | -35.1188 | -29.7377
16.7928 | -35.1194 | -29.746
16.7989 | -35.12 | -29.7541
16.805 | -35.1206 | -29.7625
16.8111 | -35.1213 | -29.7708
16.8172 | -35.1219 | -29.7794
16.8233 | -35.1225 | -29.788
16.8294 | -35.1232 | -29.7964
16.8356 | -35.1238 | -29.8049
16.8417 | -35.1244 | -29.8132
16.8478 | -35.125 | -29.8214
16.8539 | -35.1256 | -29.8297
16.86 | -35.1262 | -29.8379
16.8661 | -35.1268 | -29.8462
16.8722 | -35.1275 | -29.8545
16.8783 | -35.1282 | -29.8625
16.8844 | -35.1288 | -29.8708
16.8905 | -35.1295 | -29.8792
16.8966 | -35.1301 | -29.8875
16.9027 | -35.1307 | -29.8956
16.9088 | -35.1314 | -29.9043
16.9149 | -35.132 | -29.9128
16.921 | -35.1327 | -29.9211
16.9271 | -35.1333 | -29.9298
16.9332 | -35.1339 | -29.938
16.9393 | -35.1345 | -29.9465
16.9454 | -35.1351 | -29.9545
16.9515 | -35.1357 | -29.963
16.9576 | -35.1363 | -29.971
16.9637 | -35.137 | -29.9793
16.9698 | -35.1377 | -29.9873
16.9759 | -35.1383 | -29.9955
16.9821 | -35.139 | -30.0039
16.9882 | -35.1396 | -30.0125
16.9943 | -35.1402 | -30.021
17.0004 | -35.1409 | -30.0294
17.0065 | -35.1415 | -30.0377
17.0126 | -35.1421 | -30.046
17.0187 | -35.1427 | -30.0544
17.0248 | -35.1433 | -30.0623
17.0309 | -35.144 | -30.0707
17.037 | -35.1446 | -30.0791
17.0431 | -35.1453 | -30.0872
17.0492 | -35.1459 | -30.0953
17.0553 | -35.1464 | -30.1037
17.0614 | -35.1471 | -30.1119
17.0675 | -35.1477 | -30.1205
17.0736 | -35.1483 | -30.1288
17.0797 | -35.149 | -30.1367
17.0858 | -35.1496 | -30.1451
17.0919 | -35.1502 | -30.1536
17.098 | -35.1508 | -30.1617
17.1041 | -35.1514 | -30.1698
17.1102 | -35.1521 | -30.1784
17.1163 | -35.1528 | -30.1868
17.1225 | -35.1534 | -30.1949
17.1286 | -35.154 | -30.2029
17.1347 | -35.1546 | -30.2111
17.1408 | -35.1553 | -30.2193
17.1469 | -35.1559 | -30.2274
17.153 | -35.1566 | -30.2355
17.1591 | -35.1572 | -30.2436
17.1652 | -35.1578 | -30.2515
17.1713 | -35.1584 | -30.2595
17.1774 | -35.1591 | -30.2679
17.1835 | -35.1597 | -30.2763
17.1896 | -35.1603 | -30.2843
17.1957 | -35.1609 | -30.2927
17.2018 | -35.1615 | -30.3012
17.2079 | -35.1621 | -30.3093
17.214 | -35.1627 | -30.3175
17.2201 | -35.1633 | -30.3258
17.2262 | -35.164 | -30.3341
17.2323 | -35.1646 | -30.3424
17.2384 | -35.1652 | -30.3506
17.2445 | -35.1658 | -30.3586
17.2506 | -35.1664 | -30.3669
17.2567 | -35.167 | -30.3751
17.2628 | -35.1676 | -30.3837
17.269 | -35.1682 | -30.392
17.2751 | -35.1689 | -30.4002
17.2812 | -35.1695 | -30.4083
17.2873 | -35.1701 | -30.4166
17.2934 | -35.1707 | -30.4246
17.2995 | -35.1713 | -30.4327
17.3056 | -35.1719 | -30.4409
17.3117 | -35.1725 | -30.4491
17.3178 | -35.1731 | -30.4573
17.3239 | -35.1737 | -30.4656
17.33 | -35.1743 | -30.4738
17.3361 | -35.1749 | -30.482
17.3422 | -35.1755 | -30.4901
17.3483 | -35.1762 | -30.4984
17.3544 | -35.1768 | -30.5064
17.3605 | -35.1774 | -30.5147
17.3666 | -35.1781 | -30.5232
17.3727 | -35.1787 | -30.5313
17.3788 | -35.1793 | -30.5393
17.3849 | -35.1799 | -30.5475
17.391 | -35.1805 | -30.5555
17.3971 | -35.1812 | -30.5637
17.4032 | -35.1818 | -30.5723
17.4094 | -35.1824 | -30.5806
17.4155 | -35.183 | -30.5887
17.4216 | -35.1836 | -30.5967
17.4277 | -35.1842 | -30.6052
17.4338 | -35.1849 | -30.6133
17.4399 | -35.1855 | -30.6216
17.446 | -35.1861 | -30.6294
17.4521 | -35.1867 | -30.6375
17.4582 | -35.1873 | -30.6456
17.4643 | -35.1879 | -30.6541
17.4704 | -35.1885 | -30.6621
17.4765 | -35.1891 | -30.6704
17.4826 | -35.1897 | -30.6787
17.4887 | -35.1904 | -30.687
17.4948 | -35.191 | -30.6951
17.5009 | -35.1915 | -30.703
17.507 | -35.1922 | -30.7115
17.5131 | -35.1928 | -30.7196
17.5192 | -35.1934 | -30.7277
17.5253 | -35.1939 | -30.7359
17.5314 | -35.1945 | -30.7442
17.5375 | -35.1951 | -30.7525
17.5436 | -35.1957 | -30.7608
17.5497 | -35.1963 | -30.7692
17.5559 | -35.1969 | -30.7771
17.562 | -35.1975 | -30.7852
17.5681 | -35.1981 | -30.7933
17.5742 | -35.1987 | -30.8017
17.5803 | -35.1993 | -30.8097
17.5864 | -35.2 | -30.8175
17.5925 | -35.2006 | -30.826
17.5986 | -35.2012 | -30.834
17.6047 | -35.2019 | -30.8426
17.6108 | -35.2025 | -30.8506
17.6169 | -35.2031 | -30.8586
17.623 | -35.2037 | -30.8666
17.6291 | -35.2043 | -30.8747
17.6352 | -35.2049 | -30.8827
17.6413 | -35.2055 | -30.891
17.6474 | -35.2061 | -30.8991
17.6535 | -35.2067 | -30.9073
17.6596 | -35.2073 | -30.9155
17.6657 | -35.2079 | -30.9234
17.6718 | -35.2085 | -30.9317
17.6779 | -35.2091 | -30.9399
17.684 | -35.2097 | -30.948
17.6901 | -35.2104 | -30.9561
17.6963 | -35.2109 | -30.9639
17.7024 | -35.2116 | -30.972
17.7085 | -35.2121 | -30.98
17.7146 | -35.2128 | -30.9879
17.7207 | -35.2134 | -30.9961
17.7268 | -35.214 | -31.0041
17.7329 | -35.2146 | -31.0125
17.739 | -35.2152 | -31.0209
17.7451 | -35.2158 | -31.029
17.7512 | -35.2164 | -31.0373
17.7573 | -35.217 | -31.0456
17.7634 | -35.2176 | -31.0535
17.7695 | -35.2182 | -31.0617
17.7756 | -35.2188 | -31.0699
17.7817 | -35.2194 | -31.0778
17.7878 | -35.22 | -31.0859
17.7939 | -35.2206 | -31.0939
17.8 | -35.2213 | -31.1018
17.8061 | -35.2219 | -31.11
17.8122 | -35.2225 | -31.1181
17.8183 | -35.2231 | -31.1263
17.8244 | -35.2237 | -31.1345
17.8305 | -35.2243 | -31.1427
17.8366 | -35.225 | -31.1504
17.8428 | -35.2256 | -31.1587
17.8489 | -35.2262 | -31.1667
17.855 | -35.2268 | -31.1746
17.8611 | -35.2274 | -31.1825
17.8672 | -35.228 | -31.1906
17.8733 | -35.2286 | -31.1989
17.8794 | -35.2293 | -31.2072
17.8855 | -35.2299 | -31.2152
17.8916 | -35.2305 | -31.2231
17.8977 | -35.2311 | -31.2312
17.9038 | -35.2317 | -31.2391
17.9099 | -35.2323 | -31.2473
17.916 | -35.2329 | -31.2554
17.9221 | -35.2335 | -31.2635
17.9282 | -35.2341 | -31.2716
17.9343 | -35.2347 | -31.2797
17.9404 | -35.2353 | -31.2877
17.9465 | -35.2359 | -31.2961
17.9526 | -35.2365 | -31.3043
17.9587 | -35.2371 | -31.3123
17.9648 | -35.2377 | -31.3204
17.9709 | -35.2383 | -31.3284
17.977 | -35.2389 | -31.3362
17.9832 | -35.2395 | -31.3444
17.9893 | -35.2401 | -31.3526
17.9954 | -35.2407 | -31.3608
18.0015 | -35.2414 | -31.369
18.0076 | -35.2419 | -31.3765
18.0137 | -35.2425 | -31.3851
18.0198 | -35.2431 | -31.3932
18.0259 | -35.2436 | -31.4014
18.032 | -35.2442 | -31.4093
18.0381 | -35.2448 | -31.4173
18.0442 | -35.2454 | -31.4256
18.0503 | -35.246 | -31.4339
18.0564 | -35.2466 | -31.4419
18.0625 | -35.2472 | -31.45
18.0686 | -35.2478 | -31.4579
18.0747 | -35.2484 | -31.466
18.0808 | -35.249 | -31.4738
18.0869 | -35.2496 | -31.4818
18.093 | -35.2502 | -31.4899
18.0991 | -35.2508 | -31.4983
18.1052 | -35.2514 | -31.5068
18.1113 | -35.252 | -31.5145
18.1174 | -35.2526 | -31.5223
18.1236 | -35.2533 | -31.5305
18.1297 | -35.2538 | -31.5384
18.1358 | -35.2545 | -31.5465
18.1419 | -35.2551 | -31.5547
18.148 | -35.2558 | -31.5628
18.1541 | -35.2564 | -31.5707
18.1602 | -35.257 | -31.5788
18.1663 | -35.2577 | -31.5869
18.1724 | -35.2583 | -31.5951
18.1785 | -35.2589 | -31.6032
18.1846 | -35.2595 | -31.6112
18.1907 | -35.26 | -31.6196
18.1968 | -35.2606 | -31.6278
18.2029 | -35.2612 | -31.6354
18.209 | -35.2618 | -31.6437
18.2151 | -35.2624 | -31.652
18.2212 | -35.2631 | -31.6598
18.2273 | -35.2637 | -31.6679
18.2334 | -35.2643 | -31.6759
18.2395 | -35.2649 | -31.6838
18.2456 | -35.2655 | -31.6917
18.2517 | -35.2661 | -31.6998
18.2578 | -35.2666 | -31.7075
18.2639 | -35.2673 | -31.7158
18.2701 | -35.2679 | -31.7238
18.2762 | -35.2685 | -31.732
18.2823 | -35.2691 | -31.7401
18.2884 | -35.2697 | -31.7483
18.2945 | -35.2704 | -31.7561
18.3006 | -35.271 | -31.7642
18.3067 | -35.2716 | -31.7721
18.3128 | -35.2722 | -31.7801
18.3189 | -35.2728 | -31.7883
18.325 | -35.2734 | -31.7959
18.3311 | -35.274 | -31.8041
18.3372 | -35.2746 | -31.8117
18.3433 | -35.2752 | -31.8195
18.3494 | -35.2758 | -31.8276
18.3555 | -35.2764 | -31.8352
18.3616 | -35.277 | -31.8429
18.3677 | -35.2777 | -31.8507
18.3738 | -35.2782 | -31.8586
18.3799 | -35.2788 | -31.8663
18.386 | -35.2794 | -31.8741
18.3921 | -35.28 | -31.882
18.3982 | -35.2806 | -31.8897
18.4043 | -35.2812 | -31.8976
18.4105 | -35.2818 | -31.9055
18.4166 | -35.2824 | -31.9132
18.4227 | -35.283 | -31.9213
18.4288 | -35.2836 | -31.9293
18.4349 | -35.2842 | -31.9372
18.441 | -35.2848 | -31.9452
18.4471 | -35.2854 | -31.9532
18.4532 | -35.286 | -31.9613
18.4593 | -35.2866 | -31.9692
18.4654 | -35.2872 | -31.9769
18.4715 | -35.2878 | -31.985
18.4776 | -35.2884 | -31.9932
18.4837 | -35.289 | -32.0012
18.4898 | -35.2896 | -32.0092
18.4959 | -35.2903 | -32.0172
18.502 | -35.2908 | -32.0252
18.5081 | -35.2915 | -32.0333
18.5142 | -35.292 | -32.0416
18.5203 | -35.2926 | -32.0492
18.5264 | -35.2932 | -32.0569
18.5325 | -35.2938 | -32.065
18.5386 | -35.2945 | -32.073
18.5447 | -35.2951 | -32.0808
18.5508 | -35.2957 | -32.0885
18.557 | -35.2963 | -32.0968
18.5631 | -35.2969 | -32.1047
18.5692 | -35.2975 | -32.1125
18.5753 | -35.2981 | -32.1201
18.5814 | -35.2987 | -32.1282
18.5875 | -35.2993 | -32.1364
18.5936 | -35.2999 | -32.1446
18.5997 | -35.3005 | -32.1526
18.6058 | -35.301 | -32.1604
18.6119 | -35.3016 | -32.1683
18.618 | -35.3023 | -32.1763
18.6241 | -35.3029 | -32.1844
18.6302 | -35.3036 | -32.1927
18.6363 | -35.3042 | -32.2004
18.6424 | -35.3048 | -32.2081
18.6485 | -35.3053 | -32.2161
18.6546 | -35.3059 | -32.224
18.6607 | -35.3065 | -32.2317
18.6668 | -35.3072 | -32.2396
18.6729 | -35.3077 | -32.2476
18.679 | -35.3084 | -32.2553
18.6851 | -35.309 | -32.263
18.6912 | -35.3096 | -32.271
18.6974 | -35.3102 | -32.2789
18.7035 | -35.3108 | -32.2868
18.7096 | -35.3114 | -32.2948
18.7157 | -35.312 | -32.3028
18.7218 | -35.3126 | -32.3107
18.7279 | -35.3133 | -32.3184
18.734 | -35.3139 | -32.3263
18.7401 | -35.3145 | -32.3341
18.7462 | -35.3151 | -32.3414
18.7523 | -35.3158 | -32.3493
18.7584 | -35.3164 | -32.3573
18.7645 | -35.317 | -32.3653
18.7706 | -35.3176 | -32.3733
18.7767 | -35.3182 | -32.3811
18.7828 | -35.3188 | -32.3883
18.7889 | -35.3193 | -32.3961
18.795 | -35.32 | -32.404
18.8011 | -35.3206 | -32.412
18.8072 | -35.3212 | -32.4199
18.8133 | -35.3218 | -32.4278
18.8194 | -35.3224 | -32.4353
18.8255 | -35.323 | -32.443
18.8316 | -35.3236 | -32.4512
18.8377 | -35.3242 | -32.4588
18.8439 | -35.3248 | -32.4666
18.85 | -35.3254 | -32.4745
18.8561 | -35.326 | -32.4821
18.8622 | -35.3266 | -32.4901
18.8683 | -35.3272 | -32.4978
18.8744 | -35.3278 | -32.5054
18.8805 | -35.3284 | -32.5133
18.8866 | -35.329 | -32.5213
18.8927 | -35.3297 | -32.5289
18.8988 | -35.3302 | -32.5367
18.9049 | -35.3309 | -32.5445
18.911 | -35.3315 | -32.5523
18.9171 | -35.3321 | -32.5602
18.9232 | -35.3328 | -32.5676
18.9293 | -35.3333 | -32.5754
18.9354 | -35.3339 | -32.583
18.9415 | -35.3345 | -32.5908
18.9476 | -35.3351 | -32.5988
18.9537 | -35.3358 | -32.6068
18.9598 | -35.3364 | -32.6139
18.9659 | -35.337 | -32.6215
18.972 | -35.3375 | -32.6292
18.9781 | -35.3381 | -32.6369
18.9843 | -35.3388 | -32.6447
18.9904 | -35.3393 | -32.6527
18.9965 | -35.34 | -32.6602
19.0026 | -35.3406 | -32.6682
19.0087 | -35.3412 | -32.676
19.0148 | -35.3418 | -32.6836
19.0209 | -35.3424 | -32.6917
19.027 | -35.343 | -32.6998
19.0331 | -35.3436 | -32.7076
19.0392 | -35.3443 | -32.7153
19.0453 | -35.3449 | -32.723
19.0514 | -35.3455 | -32.7305
19.0575 | -35.3461 | -32.7384
19.0636 | -35.3467 | -32.7455
19.0697 | -35.3473 | -32.7536
19.0758 | -35.348 | -32.7612
19.0819 | -35.3486 | -32.7692
19.088 | -35.3492 | -32.7766
19.0941 | -35.3498 | -32.7841
19.1002 | -35.3504 | -32.7917
19.1063 | -35.351 | -32.7996
19.1124 | -35.3517 | -32.8073
19.1185 | -35.3523 | -32.8153
19.1246 | -35.3528 | -32.8228
19.1308 | -35.3535 | -32.8301
19.1369 | -35.3541 | -32.8375
19.143 | -35.3546 | -32.8453
19.1491 | -35.3553 | -32.8531
19.1552 | -35.3558 | -32.8613
19.1613 | -35.3564 | -32.869
19.1674 | -35.357 | -32.877
19.1735 | -35.3576 | -32.8848
19.1796 | -35.3583 | -32.8923
19.1857 | -35.3589 | -32.9001
19.1918 | -35.3595 | -32.908
19.1979 | -35.3601 | -32.9157
19.204 | -35.3607 | -32.9234
19.2101 | -35.3613 | -32.9309
19.2162 | -35.3619 | -32.939
19.2223 | -35.3625 | -32.9463
19.2284 | -35.3631 | -32.9544
19.2345 | -35.3637 | -32.9623
19.2406 | -35.3643 | -32.9699
19.2467 | -35.3649 | -32.9777
19.2528 | -35.3655 | -32.9855
19.2589 | -35.3661 | -32.9932
19.265 | -35.3667 | -33.001
19.2712 | -35.3673 | -33.0088
19.2773 | -35.3678 | -33.0161
19.2834 | -35.3684 | -33.0236
19.2895 | -35.3691 | -33.0313
19.2956 | -35.3697 | -33.0393
19.3017 | -35.3703 | -33.0468
19.3078 | -35.3709 | -33.0544
19.3139 | -35.3716 | -33.0622
19.32 | -35.3722 | -33.0698
19.3261 | -35.3728 | -33.0775
19.3322 | -35.3734 | -33.0851
19.3383 | -35.3741 | -33.0925
19.3444 | -35.3747 | -33.1
19.3505 | -35.3753 | -33.1075
19.3566 | -35.3758 | -33.1152
19.3627 | -35.3765 | -33.123
19.3688 | -35.3771 | -33.1309
19.3749 | -35.3777 | -33.1389
19.381 | -35.3783 | -33.146
19.3871 | -35.3789 | -33.1538
19.3932 | -35.3794 | -33.1619
19.3993 | -35.38 | -33.1696
19.4054 | -35.3806 | -33.1775
19.4115 | -35.3812 | -33.1852
19.4177 | -35.3818 | -33.1925
19.4238 | -35.3825 | -33.2
19.4299 | -35.3831 | -33.2074
19.436 | -35.3837 | -33.2153
19.4421 | -35.3843 | -33.2227
19.4482 | -35.3849 | -33.2305
19.4543 | -35.3855 | -33.2374
19.4604 | -35.3861 | -33.2449
19.4665 | -35.3867 | -33.2525
19.4726 | -35.3874 | -33.2597
19.4787 | -35.388 | -33.2669
19.4848 | -35.3885 | -33.2745
19.4909 | -35.3892 | -33.2821
19.497 | -35.3897 | -33.2897
19.5031 | -35.3902 | -33.2972
19.5092 | -35.3908 | -33.305
19.5153 | -35.3914 | -33.3128
19.5214 | -35.392 | -33.3204
19.5275 | -35.3925 | -33.3281
19.5336 | -35.3931 | -33.3357
19.5397 | -35.3938 | -33.3438
19.5458 | -35.3944 | -33.3512
19.5519 | -35.395 | -33.3589
19.5581 | -35.3955 | -33.3663
19.5642 | -35.3962 | -33.3736
19.5703 | -35.3967 | -33.381
19.5764 | -35.3973 | -33.3882
19.5825 | -35.3979 | -33.3959
19.5886 | -35.3985 | -33.4035
19.5947 | -35.3992 | -33.4106
19.6008 | -35.3998 | -33.4185
19.6069 | -35.4004 | -33.4259
19.613 | -35.401 | -33.434
19.6191 | -35.4016 | -33.4417
19.6252 | -35.4021 | -33.4494
19.6313 | -35.4028 | -33.4565
19.6374 | -35.4033 | -33.4644
19.6435 | -35.4039 | -33.4717
19.6496 | -35.4045 | -33.4789
19.6557 | -35.4051 | -33.4867
19.6618 | -35.4057 | -33.4942
19.6679 | -35.4063 | -33.5019
19.674 | -35.4069 | -33.5098
19.6801 | -35.4075 | -33.5173
19.6862 | -35.4081 | -33.5248
19.6923 | -35.4087 | -33.532
19.6984 | -35.4094 | -33.5393
19.7046 | -35.41 | -33.547
19.7107 | -35.4106 | -33.5544
19.7168 | -35.4112 | -33.5621
19.7229 | -35.4117 | -33.5695
19.729 | -35.4123 | -33.5769
19.7351 | -35.4129 | -33.5841
19.7412 | -35.4135 | -33.5913
19.7473 | -35.4141 | -33.599
19.7534 | -35.4147 | -33.6066
19.7595 | -35.4153 | -33.6142
19.7656 | -35.4159 | -33.6221
19.7717 | -35.4164 | -33.6292
19.7778 | -35.417 | -33.6369
19.7839 | -35.4176 | -33.6443
19.79 | -35.4182 | -33.6517
19.7961 | -35.4188 | -33.6594
19.8022 | -35.4193 | -33.6669
19.8083 | -35.4199 | -33.6745
19.8144 | -35.4205 | -33.6818
19.8205 | -35.4211 | -33.6893
19.8266 | -35.4217 | -33.6968
19.8327 | -35.4222 | -33.7043
19.8388 | -35.4228 | -33.7119
19.845 | -35.4234 | -33.7194
19.8511 | -35.4239 | -33.727
19.8572 | -35.4245 | -33.7347
19.8633 | -35.4251 | -33.7424
19.8694 | -35.4257 | -33.75
19.8755 | -35.4264 | -33.7572
19.8816 | -35.427 | -33.7647
19.8877 | -35.4276 | -33.7723
19.8938 | -35.4281 | -33.7796
19.8999 | -35.4287 | -33.7871
19.906 | -35.4293 | -33.795
19.9121 | -35.43 | -33.8022
19.9182 | -35.4306 | -33.8095
19.9243 | -35.4311 | -33.8173
19.9304 | -35.4317 | -33.8251
19.9365 | -35.4323 | -33.8328
19.9426 | -35.4329 | -33.8398
19.9487 | -35.4335 | -33.8471
19.9548 | -35.4341 | -33.8548
19.9609 | -35.4348 | -33.8622
19.967 | -35.4354 | -33.8694
19.9731 | -35.436 | -33.8772
19.9792 | -35.4365 | -33.8845
19.9853 | -35.4371 | -33.8922
19.9915 | -35.4378 | -33.8997
19.9976 | -35.4384 | -33.9071
20.0037 | -35.4389 | -33.9146
20.0098 | -35.4395 | -33.9219
20.0159 | -35.4401 | -33.9291
20.022 | -35.4407 | -33.9367
20.0281 | -35.4414 | -33.944
20.0342 | -35.4419 | -33.9511
20.0403 | -35.4426 | -33.958
20.0464 | -35.4432 | -33.966
20.0525 | -35.4438 | -33.9733
20.0586 | -35.4444 | -33.9809
20.0647 | -35.445 | -33.9886
20.0708 | -35.4456 | -33.9961
20.0769 | -35.4461 | -34.0036
20.083 | -35.4467 | -34.0108
20.0891 | -35.4473 | -34.018
20.0952 | -35.4479 | -34.025
20.1013 | -35.4484 | -34.0321
20.1074 | -35.449 | -34.0399
20.1135 | -35.4496 | -34.047
20.1196 | -35.4502 | -34.0541
20.1257 | -35.4508 | -34.0617
20.1319 | -35.4514 | -34.0693
20.138 | -35.452 | -34.0767
20.1441 | -35.4525 | -34.0842
20.1502 | -35.4531 | -34.0919
20.1563 | -35.4537 | -34.0993
20.1624 | -35.4543 | -34.1066
20.1685 | -35.4549 | -34.1138
20.1746 | -35.4555 | -34.1215
20.1807 | -35.4561 | -34.1289
20.1868 | -35.4567 | -34.1364
20.1929 | -35.4573 | -34.1433
20.199 | -35.4578 | -34.1508
20.2051 | -35.4584 | -34.1581
20.2112 | -35.4591 | -34.1652
20.2173 | -35.4596 | -34.1729
20.2234 | -35.4602 | -34.1805
20.2295 | -35.4607 | -34.1875
20.2356 | -35.4613 | -34.1954
20.2417 | -35.4619 | -34.2025
20.2478 | -35.4625 | -34.2096
20.2539 | -35.463 | -34.2174
20.26 | -35.4636 | -34.2244
20.2661 | -35.4641 | -34.2319
20.2723 | -35.4647 | -34.239
20.2784 | -35.4653 | -34.2466
20.2845 | -35.4659 | -34.254
20.2906 | -35.4665 | -34.2613
20.2967 | -35.4671 | -34.2684
20.3028 | -35.4677 | -34.2755
20.3089 | -35.4683 | -34.2831
20.315 | -35.4689 | -34.2905
20.3211 | -35.4694 | -34.2979
20.3272 | -35.47 | -34.3053
20.3333 | -35.4705 | -34.312
20.3394 | -35.4711 | -34.3192
20.3455 | -35.4717 | -34.3265
20.3516 | -35.4723 | -34.3337
20.3577 | -35.4728 | -34.3408
20.3638 | -35.4735 | -34.3482
20.3699 | -35.4741 | -34.3554
20.376 | -35.4746 | -34.3628
20.3821 | -35.4752 | -34.37
20.3882 | -35.4758 | -34.3773
20.3943 | -35.4764 | -34.385
20.4004 | -35.477 | -34.392
20.4065 | -35.4776 | -34.3994
20.4126 | -35.4781 | -34.4067
20.4188 | -35.4787 | -34.4146
20.4249 | -35.4793 | -34.4222
20.431 | -35.4799 | -34.4296
20.4371 | -35.4805 | -34.4364
20.4432 | -35.481 | -34.4437
20.4493 | -35.4816 | -34.4508
20.4554 | -35.4822 | -34.4583
20.4615 | -35.4828 | -34.4655
20.4676 | -35.4833 | -34.4731
20.4737 | -35.4839 | -34.4804
20.4798 | -35.4845 | -34.4879
20.4859 | -35.4851 | -34.495
20.492 | -35.4857 | -34.5021
20.4981 | -35.4862 | -34.5094
20.5042 | -35.4868 | -34.5161
20.5103 | -35.4874 | -34.5231
20.5164 | -35.488 | -34.5306
20.5225 | -35.4885 | -34.5381
20.5286 | -35.4891 | -34.5456
20.5347 | -35.4896 | -34.553
20.5408 | -35.4902 | -34.5601
20.5469 | -35.4908 | -34.5675
20.553 | -35.4914 | -34.5746
20.5592 | -35.492 | -34.5818
20.5653 | -35.4925 | -34.5892
20.5714 | -35.4932 | -34.5958
20.5775 | -35.4937 | -34.6029
20.5836 | -35.4943 | -34.6104
20.5897 | -35.4949 | -34.6176
20.5958 | -35.4955 | -34.6245
20.6019 | -35.496 | -34.6316
20.608 | -35.4965 | -34.639
20.6141 | -35.4971 | -34.646
20.6202 | -35.4977 | -34.6532
20.6263 | -35.4982 | -34.66
20.6324 | -35.4988 | -34.667
20.6385 | -35.4994 | -34.6743
20.6446 | -35.4999 | -34.6819
20.6507 | -35.5005 | -34.6883
20.6568 | -35.5011 | -34.6953
20.6629 | -35.5017 | -34.7026
20.669 | -35.5022 | -34.7098
20.6751 | -35.5028 | -34.7171
20.6812 | -35.5033 | -34.7239
20.6873 | -35.5039 | -34.7315
20.6934 | -35.5045 | -34.7385
20.6995 | -35.5051 | -34.7458
20.7057 | -35.5057 | -34.7531
20.7118 | -35.5063 | -34.7601
20.7179 | -35.5069 | -34.7665
20.724 | -35.5074 | -34.7735
20.7301 | -35.508 | -34.781
20.7362 | -35.5086 | -34.7884
20.7423 | -35.5091 | -34.795
20.7484 | -35.5097 | -34.8022
20.7545 | -35.5103 | -34.8094
20.7606 | -35.5109 | -34.8166
20.7667 | -35.5115 | -34.8239
20.7728 | -35.5122 | -34.8309
20.7789 | -35.5127 | -34.8381
20.785 | -35.5133 | -34.8452
20.7911 | -35.5139 | -34.8525
20.7972 | -35.5144 | -34.8597
20.8033 | -35.5149 | -34.8666
20.8094 | -35.5155 | -34.8741
20.8155 | -35.516 | -34.8813
20.8216 | -35.5166 | -34.8886
20.8277 | -35.5172 | -34.8953
20.8338 | -35.5177 | -34.9025
20.8399 | -35.5184 | -34.9095
20.8461 | -35.5189 | -34.9168
20.8522 | -35.5195 | -34.9238
20.8583 | -35.5201 | -34.9309
20.8644 | -35.5207 | -34.9383
20.8705 | -35.5212 | -34.9458
20.8766 | -35.5218 | -34.9531
20.8827 | -35.5224 | -34.9605
20.8888 | -35.523 | -34.9679
20.8949 | -35.5235 | -34.9746
20.901 | -35.5241 | -34.9817
20.9071 | -35.5246 | -34.9884
20.9132 | -35.5252 | -34.9954
20.9193 | -35.5258 | -35.0028
20.9254 | -35.5264 | -35.0096
20.9315 | -35.5269 | -35.0169
20.9376 | -35.5275 | -35.0237
20.9437 | -35.5281 | -35.0312
20.9498 | -35.5287 | -35.0385
20.9559 | -35.5293 | -35.0452
20.962 | -35.5299 | -35.0524
20.9681 | -35.5304 | -35.0593
20.9742 | -35.531 | -35.0661
20.9803 | -35.5316 | -35.0732
20.9864 | -35.5321 | -35.0805
20.9926 | -35.5327 | -35.0878
20.9987 | -35.5333 | -35.095
21.0048 | -35.5338 | -35.1024
21.0109 | -35.5344 | -35.1092
21.017 | -35.535 | -35.1163
21.0231 | -35.5355 | -35.1233
21.0292 | -35.5361 | -35.1307
21.0353 | -35.5367 | -35.1374
21.0414 | -35.5373 | -35.1445
21.0475 | -35.5379 | -35.1518
21.0536 | -35.5385 | -35.159
21.0597 | -35.5391 | -35.1657
21.0658 | -35.5396 | -35.1728
21.0719 | -35.5402 | -35.1796
21.078 | -35.5407 | -35.1867
21.0841 | -35.5413 | -35.1936
21.0902 | -35.5419 | -35.2008
21.0963 | -35.5424 | -35.208
21.1024 | -35.543 | -35.2152
21.1085 | -35.5435 | -35.2218
21.1146 | -35.544 | -35.2289
21.1207 | -35.5446 | -35.236
21.1268 | -35.5452 | -35.2432
21.133 | -35.5459 | -35.2503
21.1391 | -35.5464 | -35.2569
21.1452 | -35.547 | -35.2641
21.1513 | -35.5476 | -35.271
21.1574 | -35.5481 | -35.2779
21.1635 | -35.5486 | -35.2846
21.1696 | -35.5492 | -35.2914
21.1757 | -35.5498 | -35.2981
21.1818 | -35.5504 | -35.3048
21.1879 | -35.5509 | -35.3117
21.194 | -35.5515 | -35.3186
21.2001 | -35.552 | -35.3256
21.2062 | -35.5526 | -35.3323
21.2123 | -35.5531 | -35.3389
21.2184 | -35.5537 | -35.3459
21.2245 | -35.5542 | -35.3531
21.2306 | -35.5548 | -35.3599
21.2367 | -35.5554 | -35.3668
21.2428 | -35.5559 | -35.3737
21.2489 | -35.5565 | -35.38
21.255 | -35.5571 | -35.387
21.2611 | -35.5577 | -35.3937
21.2672 | -35.5582 | -35.4008
21.2733 | -35.5587 | -35.4077
21.2795 | -35.5593 | -35.4143
21.2856 | -35.5599 | -35.4213
21.2917 | -35.5605 | -35.4281
21.2978 | -35.5611 | -35.4352
21.3039 | -35.5616 | -35.4423
21.31 | -35.5621 | -35.4491
21.3161 | -35.5627 | -35.4555
21.3222 | -35.5632 | -35.4622
21.3283 | -35.5638 | -35.4693
21.3344 | -35.5643 | -35.4762
21.3405 | -35.5648 | -35.4832
21.3466 | -35.5655 | -35.4901
21.3527 | -35.566 | -35.497
21.3588 | -35.5667 | -35.5037
21.3649 | -35.5672 | -35.5104
21.371 | -35.5678 | -35.517
21.3771 | -35.5684 | -35.5239
21.3832 | -35.5689 | -35.5303
21.3893 | -35.5694 | -35.5368
21.3954 | -35.5699 | -35.5434
21.4015 | -35.5705 | -35.55
21.4076 | -35.5711 | -35.5571
21.4137 | -35.5717 | -35.564
21.4199 | -35.5722 | -35.5708
21.426 | -35.5727 | -35.578
21.4321 | -35.5733 | -35.5847
21.4382 | -35.5738 | -35.5914
21.4443 | -35.5743 | -35.5983
21.4504 | -35.5749 | -35.6053
21.4565 | -35.5754 | -35.6119
21.4626 | -35.5761 | -35.6188
21.4687 | -35.5766 | -35.626
21.4748 | -35.5772 | -35.6324
21.4809 | -35.5777 | -35.6396
21.487 | -35.5783 | -35.6467
21.4931 | -35.579 | -35.6533
21.4992 | -35.5795 | -35.6594
21.5053 | -35.5801 | -35.6664
21.5114 | -35.5807 | -35.6732
21.5175 | -35.5812 | -35.6802
21.5236 | -35.5818 | -35.6872
21.5297 | -35.5823 | -35.6937
21.5358 | -35.5828 | -35.7001
21.5419 | -35.5833 | -35.7073
21.548 | -35.5839 | -35.7141
21.5541 | -35.5844 | -35.7213
21.5602 | -35.5849 | -35.7281
21.5664 | -35.5855 | -35.7343
21.5725 | -35.586 | -35.7413
21.5786 | -35.5865 | -35.748
21.5847 | -35.587 | -35.7549
21.5908 | -35.5876 | -35.7615
21.5969 | -35.5881 | -35.7679
21.603 | -35.5887 | -35.7747
21.6091 | -35.5892 | -35.7814
21.6152 | -35.5899 | -35.7883
21.6213 | -35.5903 | -35.7952
21.6274 | -35.591 | -35.8019
21.6335 | -35.5915 | -35.8087
21.6396 | -35.592 | -35.8158
21.6457 | -35.5926 | -35.8228
21.6518 | -35.5932 | -35.8293
21.6579 | -35.5937 | -35.8357
21.664 | -35.5942 | -35.8429
21.6701 | -35.5948 | -35.8492
21.6762 | -35.5953 | -35.8561
21.6823 | -35.5959 | -35.8627
21.6884 | -35.5964 | -35.8693
21.6945 | -35.597 | -35.8759
21.7006 | -35.5975 | -35.8828
21.7068 | -35.598 | -35.8894
21.7129 | -35.5985 | -35.8961
21.719 | -35.5991 | -35.9031
21.7251 | -35.5996 | -35.9101
21.7312 | -35.6002 | -35.9164
21.7373 | -35.6009 | -35.9231
21.7434 | -35.6014 | -35.93
21.7495 | -35.6019 | -35.9367
21.7556 | -35.6025 | -35.9437
21.7617 | -35.603 | -35.9503
21.7678 | -35.6036 | -35.9574
21.7739 | -35.6041 | -35.964
21.78 | -35.6047 | -35.9708
21.7861 | -35.6053 | -35.9775
21.7922 | -35.6058 | -35.9843
21.7983 | -35.6063 | -35.9914
21.8044 | -35.607 | -35.9986
21.8105 | -35.6074 | -36.0051
21.8166 | -35.608 | -36.0119
21.8227 | -35.6085 | -36.0186
21.8288 | -35.6091 | -36.0253
21.8349 | -35.6096 | -36.0321
21.841 | -35.6103 | -36.039
21.8471 | -35.6108 | -36.0458
21.8533 | -35.6113 | -36.0525
21.8594 | -35.6119 | -36.0595
21.8655 | -35.6125 | -36.0666
21.8716 | -35.6131 | -36.0738
21.8777 | -35.6137 | -36.0799
21.8838 | -35.6142 | -36.0864
21.8899 | -35.6147 | -36.093
21.896 | -35.6153 | -36.0999
21.9021 | -35.6158 | -36.1072
21.9082 | -35.6164 | -36.114
21.9143 | -35.617 | -36.1207
21.9204 | -35.6175 | -36.1273
21.9265 | -35.618 | -36.134
21.9326 | -35.6186 | -36.1408
21.9387 | -35.6191 | -36.1475
21.9448 | -35.6197 | -36.1543
21.9509 | -35.6203 | -36.1606
21.957 | -35.6208 | -36.1672
21.9631 | -35.6214 | -36.1741
21.9692 | -35.6219 | -36.181
21.9753 | -35.6224 | -36.1875
21.9814 | -35.623 | -36.1941
21.9875 | -35.6235 | -36.2006
21.9937 | -35.624 | -36.2076
21.9998 | -35.6246 | -36.2145
22.0059 | -35.6251 | -36.2217
22.012 | -35.6257 | -36.228
22.0181 | -35.6262 | -36.2344
22.0242 | -35.6268 | -36.2411
22.0303 | -35.6273 | -36.2479
22.0364 | -35.6279 | -36.2544
22.0425 | -35.6284 | -36.261
22.0486 | -35.629 | -36.2676
22.0547 | -35.6295 | -36.274
22.0608 | -35.63 | -36.2804
22.0669 | -35.6305 | -36.2868
22.073 | -35.6311 | -36.2931
22.0791 | -35.6317 | -36.2999
22.0852 | -35.6322 | -36.3063
22.0913 | -35.6327 | -36.3131
22.0974 | -35.6332 | -36.3196
22.1035 | -35.6338 | -36.3263
22.1096 | -35.6343 | -36.3327
22.1157 | -35.6349 | -36.3394
22.1218 | -35.6354 | -36.3462
22.1279 | -35.6359 | -36.3522
22.134 | -35.6365 | -36.3587
22.1402 | -35.6371 | -36.3654
22.1463 | -35.6376 | -36.3719
22.1524 | -35.6382 | -36.3785
22.1585 | -35.6388 | -36.3854
22.1646 | -35.6393 | -36.3916
22.1707 | -35.6398 | -36.3987
22.1768 | -35.6404 | -36.4055
22.1829 | -35.6409 | -36.4121
22.189 | -35.6414 | -36.4188
22.1951 | -35.6419 | -36.4248
22.2012 | -35.6425 | -36.4313
22.2073 | -35.643 | -36.4379
22.2134 | -35.6436 | -36.4442
22.2195 | -35.6441 | -36.4509
22.2256 | -35.6447 | -36.4573
22.2317 | -35.6452 | -36.4636
22.2378 | -35.6458 | -36.4699
22.2439 | -35.6463 | -36.4768
22.25 | -35.6468 | -36.4836
22.2561 | -35.6474 | -36.4903
22.2622 | -35.6479 | -36.497
22.2683 | -35.6484 | -36.5036
22.2744 | -35.649 | -36.5102
22.2806 | -35.6495 | -36.5168
22.2867 | -35.6501 | -36.5234
22.2928 | -35.6506 | -36.5298
22.2989 | -35.6512 | -36.5366
22.305 | -35.6518 | -36.5427
22.3111 | -35.6523 | -36.5495
22.3172 | -35.6528 | -36.5561
22.3233 | -35.6534 | -36.5626
22.3294 | -35.654 | -36.5692
22.3355 | -35.6545 | -36.5754
22.3416 | -35.6551 | -36.5818
22.3477 | -35.6557 | -36.5883
22.3538 | -35.6562 | -36.5946
22.3599 | -35.6568 | -36.601
22.366 | -35.6573 | -36.6075
22.3721 | -35.6578 | -36.614
22.3782 | -35.6584 | -36.6207
22.3843 | -35.6589 | -36.6276
22.3904 | -35.6595 | -36.6342
22.3965 | -35.66 | -36.6408
22.4026 | -35.6605 | -36.6471
22.4087 | -35.661 | -36.6537
22.4148 | -35.6615 | -36.6601
22.4209 | -35.6621 | -36.6667
22.4271 | -35.6626 | -36.6731
22.4332 | -35.6631 | -36.6791
22.4393 | -35.6637 | -36.6856
22.4454 | -35.6643 | -36.6916
22.4515 | -35.6647 | -36.6982
22.4576 | -35.6653 | -36.7046
22.4637 | -35.6658 | -36.7112
22.4698 | -35.6664 | -36.7178
22.4759 | -35.667 | -36.7241
22.482 | -35.6675 | -36.7306
22.4881 | -35.6681 | -36.7372
22.4942 | -35.6686 | -36.7438
22.5003 | -35.6692 | -36.7506
22.5064 | -35.6696 | -36.757
22.5125 | -35.6701 | -36.7636
22.5186 | -35.6707 | -36.7702
22.5247 | -35.6712 | -36.7768
22.5308 | -35.6718 | -36.7834
22.5369 | -35.6723 | -36.7904
22.543 | -35.6729 | -36.7966
22.5491 | -35.6734 | -36.8027
22.5552 | -35.674 | -36.8094
22.5613 | -35.6745 | -36.8158
22.5675 | -35.675 | -36.8224
22.5736 | -35.6755 | -36.8291
22.5797 | -35.6761 | -36.8356
22.5858 | -35.6766 | -36.8419
22.5919 | -35.6772 | -36.8484
22.598 | -35.6777 | -36.8549
22.6041 | -35.6783 | -36.8612
22.6102 | -35.6788 | -36.8679
22.6163 | -35.6794 | -36.8742
22.6224 | -35.6799 | -36.8811
22.6285 | -35.6805 | -36.8875
22.6346 | -35.681 | -36.8939
22.6407 | -35.6815 | -36.9003
22.6468 | -35.682 | -36.907
22.6529 | -35.6825 | -36.9126
22.659 | -35.683 | -36.919
22.6651 | -35.6836 | -36.9253
22.6712 | -35.6841 | -36.9317
22.6773 | -35.6846 | -36.938
22.6834 | -35.6852 | -36.9444
22.6895 | -35.6857 | -36.9506
22.6956 | -35.6863 | -36.957
22.7017 | -35.6868 | -36.9634
22.7079 | -35.6873 | -36.9699
22.714 | -35.6878 | -36.9762
22.7201 | -35.6884 | -36.9829
22.7262 | -35.6889 | -36.9894
22.7323 | -35.6894 | -36.9958
22.7384 | -35.6899 | -37.002
22.7445 | -35.6905 | -37.0084
22.7506 | -35.6911 | -37.0151
22.7567 | -35.6916 | -37.0219
22.7628 | -35.6922 | -37.0279
22.7689 | -35.6928 | -37.0343
22.775 | -35.6934 | -37.0414
22.7811 | -35.694 | -37.0484
22.7872 | -35.6945 | -37.0546
22.7933 | -35.6951 | -37.0609
22.7994 | -35.6957 | -37.0674
22.8055 | -35.6962 | -37.0741
22.8116 | -35.6967 | -37.0807
22.8177 | -35.6973 | -37.0874
22.8238 | -35.6979 | -37.0943
22.8299 | -35.6983 | -37.1007
22.836 | -35.6988 | -37.1073
22.8421 | -35.6994 | -37.1136
22.8482 | -35.6999 | -37.1199
22.8544 | -35.7005 | -37.1267
22.8605 | -35.7011 | -37.1331
22.8666 | -35.7016 | -37.1395
22.8727 | -35.7021 | -37.1458
22.8788 | -35.7027 | -37.1523
22.8849 | -35.7033 | -37.1584
22.891 | -35.7038 | -37.1648
22.8971 | -35.7044 | -37.1711
22.9032 | -35.7049 | -37.1773
22.9093 | -35.7054 | -37.1838
22.9154 | -35.7059 | -37.1905
22.9215 | -35.7065 | -37.1966
22.9276 | -35.707 | -37.2032
22.9337 | -35.7075 | -37.2095
22.9398 | -35.708 | -37.2159
22.9459 | -35.7086 | -37.2226
22.952 | -35.7091 | -37.2291
22.9581 | -35.7096 | -37.2354
22.9642 | -35.7101 | -37.2422
22.9703 | -35.7106 | -37.2481
22.9764 | -35.7112 | -37.2548
22.9825 | -35.7117 | -37.2609
22.9886 | -35.7122 | -37.2673
22.9948 | -35.7127 | -37.2737
23.0009 | -35.7132 | -37.2799
23.007 | -35.7138 | -37.286
23.0131 | -35.7144 | -37.2923
23.0192 | -35.7149 | -37.2983
23.0253 | -35.7155 | -37.3049
23.0314 | -35.716 | -37.3113
23.0375 | -35.7165 | -37.318
23.0436 | -35.7171 | -37.3245
23.0497 | -35.7176 | -37.3306
23.0558 | -35.7182 | -37.337
23.0619 | -35.7188 | -37.3434
23.068 | -35.7193 | -37.3497
23.0741 | -35.7198 | -37.3558
23.0802 | -35.7203 | -37.3622
23.0863 | -35.7208 | -37.3686
23.0924 | -35.7213 | -37.375
23.0985 | -35.7218 | -37.3813
23.1046 | -35.7223 | -37.3878
23.1107 | -35.7228 | -37.3941
23.1168 | -35.7233 | -37.4006
23.1229 | -35.7239 | -37.4069
23.129 | -35.7245 | -37.4133
23.1351 | -35.725 | -37.4199
23.1413 | -35.7255 | -37.4267
23.1474 | -35.7261 | -37.4326
23.1535 | -35.7266 | -37.439
23.1596 | -35.7272 | -37.445
23.1657 | -35.7277 | -37.4514
23.1718 | -35.7283 | -37.4577
23.1779 | -35.7288 | -37.4643
23.184 | -35.7294 | -37.4703
23.1901 | -35.73 | -37.4766
23.1962 | -35.7305 | -37.4829
23.2023 | -35.731 | -37.489
23.2084 | -35.7315 | -37.4951
23.2145 | -35.732 | -37.5012
23.2206 | -35.7326 | -37.5077
23.2267 | -35.7331 | -37.5143
23.2328 | -35.7336 | -37.5205
23.2389 | -35.7341 | -37.5269
23.245 | -35.7347 | -37.5332
23.2511 | -35.7352 | -37.5394
23.2572 | -35.7357 | -37.5455
23.2633 | -35.7362 | -37.5518
23.2694 | -35.7368 | -37.5583
23.2755 | -35.7373 | -37.5647
23.2817 | -35.7378 | -37.5707
23.2878 | -35.7383 | -37.577
23.2939 | -35.7389 | -37.5833
23.3 | -35.7394 | -37.5899
23.3061 | -35.7399 | -37.5964
23.3122 | -35.7405 | -37.603
23.3183 | -35.741 | -37.6093
23.3244 | -35.7415 | -37.6156
23.3305 | -35.742 | -37.6219
23.3366 | -35.7425 | -37.6284
23.3427 | -35.743 | -37.6346
23.3488 | -35.7435 | -37.6412
23.3549 | -35.744 | -37.6474
23.361 | -35.7446 | -37.6535
23.3671 | -35.7452 | -37.6601
23.3732 | -35.7457 | -37.6662
23.3793 | -35.7462 | -37.6725
23.3854 | -35.7468 | -37.6787
23.3915 | -35.7473 | -37.6851
23.3976 | -35.7478 | -37.6912
23.4037 | -35.7483 | -37.6974
23.4098 | -35.7489 | -37.7039
23.4159 | -35.7494 | -37.7103
23.422 | -35.75 | -37.7165
23.4282 | -35.7505 | -37.7225
23.4343 | -35.7511 | -37.7287
23.4404 | -35.7515 | -37.7353
23.4465 | -35.752 | -37.7415
23.4526 | -35.7526 | -37.748
23.4587 | -35.7531 | -37.7544
23.4648 | -35.7536 | -37.761
23.4709 | -35.754 | -37.7674
23.477 | -35.7545 | -37.7739
23.4831 | -35.7551 | -37.7801
23.4892 | -35.7556 | -37.7866
23.4953 | -35.7561 | -37.7929
23.5014 | -35.7566 | -37.799
23.5075 | -35.7571 | -37.8053
23.5136 | -35.7576 | -37.8115
23.5197 | -35.7581 | -37.8179
23.5258 | -35.7586 | -37.8241
23.5319 | -35.7591 | -37.83
23.538 | -35.7596 | -37.8363
23.5441 | -35.7601 | -37.8429
23.5502 | -35.7607 | -37.8491
23.5563 | -35.7612 | -37.8552
23.5624 | -35.7617 | -37.8618
23.5686 | -35.7622 | -37.8681
23.5747 | -35.7628 | -37.8743
23.5808 | -35.7633 | -37.8807
23.5869 | -35.7638 | -37.887
23.593 | -35.7643 | -37.8932
23.5991 | -35.7649 | -37.8995
23.6052 | -35.7654 | -37.9059
23.6113 | -35.7659 | -37.9124
23.6174 | -35.7665 | -37.9183
23.6235 | -35.767 | -37.9245
23.6296 | -35.7675 | -37.9312
23.6357 | -35.7681 | -37.938
23.6418 | -35.7686 | -37.9441
23.6479 | -35.7691 | -37.9505
23.654 | -35.7697 | -37.9568
23.6601 | -35.7702 | -37.9629
23.6662 | -35.7707 | -37.9693
23.6723 | -35.7712 | -37.9758
23.6784 | -35.7717 | -37.9821
23.6845 | -35.7722 | -37.9886
23.6906 | -35.7727 | -37.9951
23.6967 | -35.7733 | -38.0016
23.7028 | -35.7738 | -38.0079
23.7089 | -35.7744 | -38.0145
23.7151 | -35.7748 | -38.0206
23.7212 | -35.7753 | -38.0271
23.7273 | -35.7759 | -38.0333
23.7334 | -35.7764 | -38.0396
23.7395 | -35.7769 | -38.0459
23.7456 | -35.7774 | -38.0525
23.7517 | -35.7779 | -38.0587
23.7578 | -35.7784 | -38.0651
23.7639 | -35.7789 | -38.0715
23.77 | -35.7795 | -38.0777
23.7761 | -35.78 | -38.0837
23.7822 | -35.7805 | -38.0904
23.7883 | -35.781 | -38.0964
23.7944 | -35.7816 | -38.1029
23.8005 | -35.7821 | -38.1094
23.8066 | -35.7827 | -38.1158
23.8127 | -35.7832 | -38.1223
23.8188 | -35.7837 | -38.1287
23.8249 | -35.7843 | -38.1353
23.831 | -35.7848 | -38.1415
23.8371 | -35.7853 | -38.1477
23.8432 | -35.7858 | -38.1541
23.8493 | -35.7863 | -38.16
23.8555 | -35.7868 | -38.1662
23.8616 | -35.7874 | -38.1728
23.8677 | -35.7879 | -38.1789
23.8738 | -35.7884 | -38.1853
23.8799 | -35.789 | -38.1915
23.886 | -35.7894 | -38.1975
23.8921 | -35.7899 | -38.2039
23.8982 | -35.7904 | -38.2104
23.9043 | -35.791 | -38.2168
23.9104 | -35.7915 | -38.2229
23.9165 | -35.792 | -38.2287
23.9226 | -35.7925 | -38.2346
23.9287 | -35.793 | -38.2406
23.9348 | -35.7935 | -38.2467
23.9409 | -35.7941 | -38.2529
23.947 | -35.7946 | -38.2592
23.9531 | -35.7951 | -38.2652
23.9592 | -35.7956 | -38.2717
23.9653 | -35.7961 | -38.2779
23.9714 | -35.7965 | -38.2841
23.9775 | -35.797 | -38.2904
23.9836 | -35.7975 | -38.2967
23.9897 | -35.7981 | -38.3031
23.9958 | -35.7986 | -38.3094
24.002 | -35.7991 | -38.3155
24.0081 | -35.7996 | -38.3216
24.0142 | -35.8001 | -38.328
24.0203 | -35.8006 | -38.3341
24.0264 | -35.801 | -38.3399
24.0325 | -35.8016 | -38.3462
24.0386 | -35.802 | -38.3524
24.0447 | -35.8025 | -38.3584
24.0508 | -35.8031 | -38.3645
24.0569 | -35.8036 | -38.3707
24.063 | -35.8041 | -38.3773
24.0691 | -35.8047 | -38.3835
24.0752 | -35.8052 | -38.3899
24.0813 | -35.8057 | -38.3962
24.0874 | -35.8062 | -38.4027
24.0935 | -35.8067 | -38.4087
24.0996 | -35.8071 | -38.4145
24.1057 | -35.8076 | -38.4204
24.1118 | -35.8081 | -38.4265
24.1179 | -35.8085 | -38.4331
24.124 | -35.809 | -38.439
24.1301 | -35.8095 | -38.4452
24.1362 | -35.81 | -38.4517
24.1424 | -35.8105 | -38.4573
24.1485 | -35.811 | -38.4633
24.1546 | -35.8115 | -38.4695
24.1607 | -35.812 | -38.4758
24.1668 | -35.8125 | -38.4818
24.1729 | -35.813 | -38.4879
24.179 | -35.8135 | -38.4941
24.1851 | -35.8141 | -38.5005
24.1912 | -35.8146 | -38.5068
24.1973 | -35.8151 | -38.5133
24.2034 | -35.8156 | -38.5192
24.2095 | -35.8161 | -38.5255
24.2156 | -35.8166 | -38.5319
24.2217 | -35.8172 | -38.5383
24.2278 | -35.8176 | -38.5445
24.2339 | -35.8181 | -38.5507
24.24 | -35.8186 | -38.5566
24.2461 | -35.8192 | -38.5628
24.2522 | -35.8197 | -38.5693
24.2583 | -35.8202 | -38.5759
24.2644 | -35.8207 | -38.5822
24.2705 | -35.8212 | -38.5888
24.2766 | -35.8217 | -38.5954
24.2827 | -35.8223 | -38.6016
24.2889 | -35.8228 | -38.6079
24.295 | -35.8233 | -38.6143
24.3011 | -35.8238 | -38.621
24.3072 | -35.8242 | -38.6267
24.3133 | -35.8247 | -38.6328
24.3194 | -35.8252 | -38.6392
24.3255 | -35.8256 | -38.6457
24.3316 | -35.8261 | -38.6518
24.3377 | -35.8266 | -38.6582
24.3438 | -35.827 | -38.6645
24.3499 | -35.8275 | -38.6707
24.356 | -35.828 | -38.677
24.3621 | -35.8285 | -38.6834
24.3682 | -35.829 | -38.6892
24.3743 | -35.8295 | -38.6955
24.3804 | -35.83 | -38.7014
24.3865 | -35.8304 | -38.7079
24.3926 | -35.8309 | -38.7145
24.3987 | -35.8314 | -38.7212
24.4048 | -35.8319 | -38.7274
24.4109 | -35.8324 | -38.7333
24.417 | -35.8328 | -38.7393
24.4231 | -35.8333 | -38.7459
24.4293 | -35.8339 | -38.7522
24.4354 | -35.8344 | -38.7581
24.4415 | -35.8348 | -38.7639
24.4476 | -35.8353 | -38.7699
24.4537 | -35.8358 | -38.7761
24.4598 | -35.8364 | -38.7825
24.4659 | -35.8368 | -38.7891
24.472 | -35.8373 | -38.7954
24.4781 | -35.8378 | -38.8018
24.4842 | -35.8382 | -38.8077
24.4903 | -35.8387 | -38.8138
24.4964 | -35.8392 | -38.8201
24.5025 | -35.8397 | -38.8266
24.5086 | -35.8402 | -38.8329
24.5147 | -35.8407 | -38.8392
24.5208 | -35.8412 | -38.8454
24.5269 | -35.8416 | -38.8513
24.533 | -35.8422 | -38.8574
24.5391 | -35.8426 | -38.8637
24.5452 | -35.8431 | -38.8698
24.5513 | -35.8436 | -38.8754
24.5574 | -35.8441 | -38.8819
24.5635 | -35.8445 | -38.8881
24.5696 | -35.845 | -38.8945
24.5758 | -35.8456 | -38.9007
24.5819 | -35.8461 | -38.9067
24.588 | -35.8465 | -38.9129
24.5941 | -35.8471 | -38.919
24.6002 | -35.8475 | -38.925
24.6063 | -35.848 | -38.9315
24.6124 | -35.8484 | -38.9377
24.6185 | -35.8489 | -38.9441
24.6246 | -35.8494 | -38.9501
24.6307 | -35.8499 | -38.9564
24.6368 | -35.8504 | -38.9624
24.6429 | -35.8509 | -38.9683
24.649 | -35.8513 | -38.9742
24.6551 | -35.8518 | -38.9804
24.6612 | -35.8523 | -38.9865
24.6673 | -35.8528 | -38.9928
24.6734 | -35.8533 | -38.999
24.6795 | -35.8537 | -39.0052
24.6856 | -35.8543 | -39.0115
24.6917 | -35.8548 | -39.0177
24.6978 | -35.8552 | -39.0235
24.7039 | -35.8557 | -39.0298
24.71 | -35.8561 | -39.0361
24.7162 | -35.8566 | -39.0423
24.7223 | -35.8571 | -39.0484
24.7284 | -35.8575 | -39.0542
24.7345 | -35.858 | -39.0605
24.7406 | -35.8585 | -39.0667
24.7467 | -35.859 | -39.073
24.7528 | -35.8594 | -39.0793
24.7589 | -35.8599 | -39.0852
24.765 | -35.8604 | -39.0912
24.7711 | -35.8608 | -39.0974
24.7772 | -35.8613 | -39.1035
24.7833 | -35.8618 | -39.1097
24.7894 | -35.8622 | -39.1158
24.7955 | -35.8627 | -39.1217
24.8016 | -35.8633 | -39.128
24.8077 | -35.8638 | -39.1342
24.8138 | -35.8643 | -39.1404
24.8199 | -35.8648 | -39.1466
24.826 | -35.8653 | -39.1529
24.8321 | -35.8657 | -39.1592
24.8382 | -35.8662 | -39.1657
24.8443 | -35.8667 | -39.1719
24.8504 | -35.8672 | -39.178
24.8565 | -35.8677 | -39.1842
24.8627 | -35.8682 | -39.1905
24.8688 | -35.8687 | -39.1965
24.8749 | -35.8692 | -39.2025
24.881 | -35.8697 | -39.2089
24.8871 | -35.8702 | -39.215
24.8932 | -35.8707 | -39.2211
24.8993 | -35.8712 | -39.2272
24.9054 | -35.8718 | -39.2335
24.9115 | -35.8723 | -39.2399
24.9176 | -35.8727 | -39.2462
24.9237 | -35.8732 | -39.2524
24.9298 | -35.8737 | -39.2584
24.9359 | -35.8742 | -39.2647
24.942 | -35.8747 | -39.2706
24.9481 | -35.8752 | -39.2767
24.9542 | -35.8756 | -39.2829
24.9603 | -35.8761 | -39.2891
24.9664 | -35.8765 | -39.2951
24.9725 | -35.877 | -39.3012
24.9786 | -35.8775 | -39.3075
24.9847 | -35.878 | -39.3139
24.9908 | -35.8784 | -39.32
24.9969 | -35.8789 | -39.326
25.0031 | -35.8794 | -39.3321
25.0092 | -35.8799 | -39.3382
25.0153 | -35.8803 | -39.3443
25.0214 | -35.8809 | -39.3509
25.0275 | -35.8814 | -39.3568
25.0336 | -35.8819 | -39.363
25.0397 | -35.8823 | -39.3692
25.0458 | -35.8828 | -39.3754
25.0519 | -35.8833 | -39.3819
25.058 | -35.8837 | -39.388
25.0641 | -35.8842 | -39.3943
25.0702 | -35.8847 | -39.4005
25.0763 | -35.8851 | -39.4067
25.0824 | -35.8855 | -39.4128
25.0885 | -35.886 | -39.4191
25.0946 | -35.8865 | -39.4252
25.1007 | -35.8869 | -39.4314
25.1068 | -35.8874 | -39.4372
25.1129 | -35.8879 | -39.4434
25.119 | -35.8883 | -39.4495
25.1251 | -35.8888 | -39.4558
25.1312 | -35.8893 | -39.462
25.1373 | -35.8897 | -39.4683
25.1435 | -35.8902 | -39.4746
25.1496 | -35.8907 | -39.4807
25.1557 | -35.8911 | -39.4868
25.1618 | -35.8916 | -39.4931
25.1679 | -35.892 | -39.4994
25.174 | -35.8925 | -39.5055
25.1801 | -35.8929 | -39.5116
25.1862 | -35.8934 | -39.5179
25.1923 | -35.8938 | -39.5243
25.1984 | -35.8943 | -39.5302
25.2045 | -35.8948 | -39.5362
25.2106 | -35.8952 | -39.5425
25.2167 | -35.8957 | -39.5489
25.2228 | -35.8961 | -39.5548
25.2289 | -35.8967 | -39.5607
25.235 | -35.8971 | -39.5667
25.2411 | -35.8976 | -39.5729
25.2472 | -35.898 | -39.5789
25.2533 | -35.8984 | -39.5853
25.2594 | -35.8989 | -39.5913
25.2655 | -35.8994 | -39.5974
25.2716 | -35.8999 | -39.6036
25.2777 | -35.9004 | -39.6101
25.2838 | -35.9008 | -39.6162
25.29 | -35.9013 | -39.6222
25.2961 | -35.9018 | -39.6282
25.3022 | -35.9022 | -39.6343
25.3083 | -35.9026 | -39.6405
25.3144 | -35.903 | -39.6465
25.3205 | -35.9034 | -39.6526
25.3266 | -35.9039 | -39.6587
25.3327 | -35.9044 | -39.6646
25.3388 | -35.9048 | -39.671
25.3449 | -35.9053 | -39.6772
25.351 | -35.9057 | -39.683
25.3571 | -35.9062 | -39.6893
25.3632 | -35.9066 | -39.6951
25.3693 | -35.9071 | -39.7011
25.3754 | -35.9075 | -39.7072
25.3815 | -35.908 | -39.7133
25.3876 | -35.9084 | -39.7195
25.3937 | -35.9089 | -39.7254
25.3998 | -35.9094 | -39.7317
25.4059 | -35.9098 | -39.7376
25.412 | -35.9103 | -39.7435
25.4181 | -35.9107 | -39.7497
25.4242 | -35.9112 | -39.7555
25.4304 | -35.9116 | -39.7615
25.4365 | -35.9121 | -39.7676
25.4426 | -35.9125 | -39.7737
25.4487 | -35.913 | -39.78
25.4548 | -35.9134 | -39.7862
25.4609 | -35.9139 | -39.7922
25.467 | -35.9143 | -39.798
25.4731 | -35.9148 | -39.8042
25.4792 | -35.9152 | -39.8104
25.4853 | -35.9157 | -39.8164
25.4914 | -35.9162 | -39.8227
25.4975 | -35.9167 | -39.8284
25.5036 | -35.9171 | -39.8343
25.5097 | -35.9176 | -39.8403
25.5158 | -35.9181 | -39.8466
25.5219 | -35.9186 | -39.8527
25.528 | -35.919 | -39.8589
25.5341 | -35.9195 | -39.8651
25.5402 | -35.9199 | -39.8713
25.5463 | -35.9204 | -39.8775
25.5524 | -35.9209 | -39.8834
25.5585 | -35.9214 | -39.8892
25.5646 | -35.9218 | -39.8953
25.5707 | -35.9222 | -39.9013
25.5769 | -35.9226 | -39.9074
25.583 | -35.9231 | -39.9136
25.5891 | -35.9236 | -39.9197
25.5952 | -35.9241 | -39.9258
25.6013 | -35.9245 | -39.9319
25.6074 | -35.925 | -39.9377
25.6135 | -35.9254 | -39.9437
25.6196 | -35.9259 | -39.9499
25.6257 | -35.9263 | -39.9561
25.6318 | -35.9267 | -39.9618
25.6379 | -35.9272 | -39.9678
25.644 | -35.9276 | -39.9737
25.6501 | -35.928 | -39.9795
25.6562 | -35.9285 | -39.9858
25.6623 | -35.929 | -39.9918
25.6684 | -35.9294 | -39.9979
25.6745 | -35.9299 | -40.0041
25.6806 | -35.9303 | -40.01
25.6867 | -35.9308 | -40.0165
25.6928 | -35.9312 | -40.0229
25.6989 | -35.9316 | -40.0289
25.705 | -35.9321 | -40.0351
25.7111 | -35.9325 | -40.0411
25.7173 | -35.9329 | -40.0471
25.7234 | -35.9334 | -40.0531
25.7295 | -35.9338 | -40.0591
25.7356 | -35.9343 | -40.0654
25.7417 | -35.9347 | -40.0713
25.7478 | -35.9351 | -40.0771
25.7539 | -35.9356 | -40.0831
25.76 | -35.9361 | -40.0894
25.7661 | -35.9365 | -40.0952
25.7722 | -35.9369 | -40.1011
25.7783 | -35.9373 | -40.1072
25.7844 | -35.9378 | -40.1132
25.7905 | -35.9382 | -40.1192
25.7966 | -35.9386 | -40.1248
25.8027 | -35.939 | -40.131
25.8088 | -35.9394 | -40.1371
25.8149 | -35.9398 | -40.1429
25.821 | -35.9402 | -40.1492
25.8271 | -35.9406 | -40.1555
25.8332 | -35.9411 | -40.1618
25.8393 | -35.9415 | -40.1677
25.8454 | -35.9419 | -40.1736
25.8515 | -35.9424 | -40.1799
25.8576 | -35.9429 | -40.1866
25.8638 | -35.9434 | -40.1931
25.8699 | -35.9438 | -40.1994
25.876 | -35.9442 | -40.2056
25.8821 | -35.9447 | -40.2116
25.8882 | -35.9452 | -40.2171
25.8943 | -35.9456 | -40.2232
25.9004 | -35.946 | -40.2292
25.9065 | -35.9465 | -40.2353
25.9126 | -35.9469 | -40.2414
25.9187 | -35.9474 | -40.2472
25.9248 | -35.9478 | -40.2533
25.9309 | -35.9482 | -40.2592
25.937 | -35.9487 | -40.2656
25.9431 | -35.9491 | -40.2715
25.9492 | -35.9495 | -40.2776
25.9553 | -35.9499 | -40.2834
25.9614 | -35.9504 | -40.2893
25.9675 | -35.9508 | -40.2953
25.9736 | -35.9512 | -40.3012
25.9797 | -35.9517 | -40.3072
25.9858 | -35.9521 | -40.3131
25.9919 | -35.9525 | -40.3194
25.998 | -35.9529 | -40.3252
26.0042 | -35.9533 | -40.3312
26.0103 | -35.9537 | -40.3373
26.0164 | -35.9542 | -40.3432
26.0225 | -35.9546 | -40.3494
26.0286 | -35.955 | -40.3554
26.0347 | -35.9554 | -40.3611
26.0408 | -35.9558 | -40.3675
26.0469 | -35.9563 | -40.3733
26.053 | -35.9567 | -40.3795
26.0591 | -35.9571 | -40.3853
26.0652 | -35.9575 | -40.3912
26.0713 | -35.9579 | -40.3972
26.0774 | -35.9583 | -40.4031
26.0835 | -35.9588 | -40.4089
26.0896 | -35.9592 | -40.4147
26.0957 | -35.9596 | -40.4205
26.1018 | -35.96 | -40.4262
26.1079 | -35.9605 | -40.4322
26.114 | -35.9608 | -40.4383
26.1201 | -35.9612 | -40.4443
26.1262 | -35.9616 | -40.4504
26.1323 | -35.9621 | -40.4563
26.1384 | -35.9625 | -40.4621
26.1445 | -35.9628 | -40.4678
26.1507 | -35.9632 | -40.4739
26.1568 | -35.9637 | -40.48
26.1629 | -35.9641 | -40.4863
26.169 | -35.9646 | -40.4922
26.1751 | -35.9649 | -40.4987
26.1812 | -35.9653 | -40.5048
26.1873 | -35.9657 | -40.5106
26.1934 | -35.9662 | -40.5165
26.1995 | -35.9667 | -40.5225
26.2056 | -35.9671 | -40.5286
26.2117 | -35.9675 | -40.5347
26.2178 | -35.9679 | -40.5409
26.2239 | -35.9684 | -40.5467
26.23 | -35.9688 | -40.5529
26.2361 | -35.9693 | -40.5586
26.2422 | -35.9697 | -40.5643
26.2483 | -35.9701 | -40.57
26.2544 | -35.9705 | -40.576
26.2605 | -35.971 | -40.5822
26.2666 | -35.9714 | -40.5884
26.2727 | -35.9718 | -40.5944
26.2788 | -35.9722 | -40.6002
26.2849 | -35.9726 | -40.6061
26.2911 | -35.973 | -40.6119
26.2972 | -35.9735 | -40.6179
26.3033 | -35.9739 | -40.6238
26.3094 | -35.9743 | -40.6301
26.3155 | -35.9747 | -40.6361
26.3216 | -35.9751 | -40.6421
26.3277 | -35.9756 | -40.6478
26.3338 | -35.976 | -40.6538
26.3399 | -35.9764 | -40.6597
26.346 | -35.9768 | -40.6651
26.3521 | -35.9772 | -40.6716
26.3582 | -35.9776 | -40.6776
26.3643 | -35.9781 | -40.6834
26.3704 | -35.9785 | -40.6892
26.3765 | -35.979 | -40.6951
26.3826 | -35.9795 | -40.7008
26.3887 | -35.9799 | -40.7069
26.3948 | -35.9802 | -40.7128
26.4009 | -35.9807 | -40.7187
26.407 | -35.9811 | -40.7248
26.4131 | -35.9816 | -40.7307
26.4192 | -35.9819 | -40.7364
26.4253 | -35.9824 | -40.7423
26.4314 | -35.9829 | -40.7482
26.4376 | -35.9833 | -40.754
26.4437 | -35.9837 | -40.7599
26.4498 | -35.9842 | -40.7661
26.4559 | -35.9845 | -40.7721
26.462 | -35.9849 | -40.7783
26.4681 | -35.9853 | -40.7843
26.4742 | -35.9858 | -40.7904
26.4803 | -35.9862 | -40.7963
26.4864 | -35.9866 | -40.8022
26.4925 | -35.9869 | -40.8082
26.4986 | -35.9874 | -40.8143
26.5047 | -35.9878 | -40.8202
26.5108 | -35.9883 | -40.8262
26.5169 | -35.9887 | -40.8321
26.523 | -35.989 | -40.838
26.5291 | -35.9894 | -40.8437
26.5352 | -35.9898 | -40.8497
26.5413 | -35.9902 | -40.8559
26.5474 | -35.9906 | -40.8617
26.5535 | -35.9911 | -40.868
26.5596 | -35.9914 | -40.874
26.5657 | -35.9919 | -40.8798
26.5718 | -35.9923 | -40.8858
26.578 | -35.9927 | -40.8919
26.5841 | -35.9931 | -40.8977
26.5902 | -35.9935 | -40.9032
26.5963 | -35.9939 | -40.9089
26.6024 | -35.9942 | -40.9152
26.6085 | -35.9946 | -40.9211
26.6146 | -35.9951 | -40.9269
26.6207 | -35.9954 | -40.933
26.6268 | -35.9959 | -40.9391
26.6329 | -35.9963 | -40.9449
26.639 | -35.9967 | -40.951
26.6451 | -35.9971 | -40.9569
26.6512 | -35.9975 | -40.9627
26.6573 | -35.9979 | -40.9687
26.6634 | -35.9982 | -40.9745
26.6695 | -35.9987 | -40.9806
26.6756 | -35.999 | -40.9868
26.6817 | -35.9994 | -40.9929
26.6878 | -35.9998 | -40.9993
26.6939 | -36.0002 | -41.0052
26.7 | -36.0006 | -41.0107
26.7061 | -36.001 | -41.0166
26.7122 | -36.0014 | -41.0225
26.7183 | -36.0018 | -41.0285
26.7245 | -36.0022 | -41.0343
26.7306 | -36.0026 | -41.0403
26.7367 | -36.003 | -41.0461
26.7428 | -36.0033 | -41.0522
26.7489 | -36.0037 | -41.0579
26.755 | -36.0041 | -41.0636
26.7611 | -36.0045 | -41.0695
26.7672 | -36.0049 | -41.0757
26.7733 | -36.0053 | -41.0813
26.7794 | -36.0058 | -41.087
26.7855 | -36.0062 | -41.0929
26.7916 | -36.0066 | -41.0985
26.7977 | -36.0071 | -41.1045
26.8038 | -36.0075 | -41.1104
26.8099 | -36.0079 | -41.1163
26.816 | -36.0083 | -41.1223
26.8221 | -36.0086 | -41.1282
26.8282 | -36.009 | -41.134
26.8343 | -36.0095 | -41.1401
26.8404 | -36.0098 | -41.1458
26.8465 | -36.0102 | -41.1516
26.8526 | -36.0107 | -41.1579
26.8587 | -36.011 | -41.1633
26.8649 | -36.0114 | -41.1691
26.871 | -36.0118 | -41.1751
26.8771 | -36.0122 | -41.1807
26.8832 | -36.0126 | -41.1867
26.8893 | -36.013 | -41.1926
26.8954 | -36.0135 | -41.1986
26.9015 | -36.0138 | -41.2046
26.9076 | -36.0142 | -41.2107
26.9137 | -36.0146 | -41.2165
26.9198 | -36.015 | -41.2223
26.9259 | -36.0154 | -41.2276
26.932 | -36.0158 | -41.2333
26.9381 | -36.0162 | -41.2393
26.9442 | -36.0166 | -41.2447
26.9503 | -36.017 | -41.2509
26.9564 | -36.0174 | -41.2571
26.9625 | -36.0178 | -41.2629
26.9686 | -36.0182 | -41.2688
26.9747 | -36.0186 | -41.2745
26.9808 | -36.019 | -41.28
26.9869 | -36.0193 | -41.2858
26.993 | -36.0197 | -41.2914
26.9991 | -36.0201 | -41.2974
27.0052 | -36.0205 | -41.3034
27.0114 | -36.0209 | -41.3094
27.0175 | -36.0212 | -41.3153
27.0236 | -36.0216 | -41.3212
27.0297 | -36.022 | -41.327
27.0358 | -36.0224 | -41.3328
27.0419 | -36.0228 | -41.3385
27.048 | -36.0232 | -41.3439
27.0541 | -36.0236 | -41.3496
27.0602 | -36.024 | -41.3558
27.0663 | -36.0244 | -41.3618
27.0724 | -36.0248 | -41.3675
27.0785 | -36.0252 | -41.3734
27.0846 | -36.0255 | -41.3791
27.0907 | -36.0259 | -41.3852
27.0968 | -36.0263 | -41.3912
27.1029 | -36.0267 | -41.3971
27.109 | -36.0271 | -41.4027
27.1151 | -36.0275 | -41.4086
27.1212 | -36.0278 | -41.4141
27.1273 | -36.0283 | -41.42
27.1334 | -36.0286 | -41.4262
27.1395 | -36.029 | -41.4319
27.1456 | -36.0294 | -41.4377
27.1518 | -36.0298 | -41.4437
27.1579 | -36.0302 | -41.4495
27.164 | -36.0306 | -41.4552
27.1701 | -36.031 | -41.4613
27.1762 | -36.0314 | -41.4674
27.1823 | -36.0317 | -41.4733
27.1884 | -36.0321 | -41.4791
27.1945 | -36.0324 | -41.4849
27.2006 | -36.0328 | -41.4905
27.2067 | -36.0332 | -41.4962
27.2128 | -36.0336 | -41.5019
27.2189 | -36.034 | -41.5077
27.225 | -36.0344 | -41.5133
27.2311 | -36.0347 | -41.5193
27.2372 | -36.0351 | -41.5252
27.2433 | -36.0355 | -41.5309
27.2494 | -36.0359 | -41.5369
27.2555 | -36.0363 | -41.543
27.2616 | -36.0367 | -41.5485
27.2677 | -36.0371 | -41.5542
27.2738 | -36.0374 | -41.56
27.2799 | -36.0379 | -41.566
27.286 | -36.0383 | -41.5717
27.2921 | -36.0387 | -41.5776
27.2983 | -36.0391 | -41.5833
27.3044 | -36.0395 | -41.5891
27.3105 | -36.0399 | -41.595
27.3166 | -36.0402 | -41.6007
27.3227 | -36.0407 | -41.6064
27.3288 | -36.0411 | -41.612
27.3349 | -36.0415 | -41.6178
27.341 | -36.0419 | -41.6236
27.3471 | -36.0423 | -41.6292
27.3532 | -36.0426 | -41.6347
27.3593 | -36.0431 | -41.6406
27.3654 | -36.0435 | -41.6464
27.3715 | -36.0439 | -41.6519
27.3776 | -36.0443 | -41.6578
27.3837 | -36.0447 | -41.6635
27.3898 | -36.0451 | -41.6693
27.3959 | -36.0455 | -41.6749
27.402 | -36.0459 | -41.6803
27.4081 | -36.0462 | -41.6862
27.4142 | -36.0466 | -41.6919
27.4203 | -36.047 | -41.6975
27.4264 | -36.0473 | -41.7032
27.4325 | -36.0477 | -41.7087
27.4387 | -36.048 | -41.7144
27.4448 | -36.0484 | -41.7201
27.4509 | -36.0488 | -41.7263
27.457 | -36.0492 | -41.732
27.4631 | -36.0496 | -41.7379
27.4692 | -36.0499 | -41.7436
27.4753 | -36.0503 | -41.7494
27.4814 | -36.0506 | -41.7548
27.4875 | -36.051 | -41.761
27.4936 | -36.0515 | -41.7668
27.4997 | -36.0519 | -41.7724
27.5058 | -36.0523 | -41.7781
27.5119 | -36.0526 | -41.7836
27.518 | -36.053 | -41.7896
27.5241 | -36.0534 | -41.7951
27.5302 | -36.0538 | -41.8012
27.5363 | -36.0542 | -41.8073
27.5424 | -36.0546 | -41.8125
27.5485 | -36.055 | -41.8183
27.5546 | -36.0554 | -41.8238
27.5607 | -36.0558 | -41.8294
27.5668 | -36.0561 | -41.835
27.5729 | -36.0565 | -41.8406
27.5791 | -36.0569 | -41.8463
27.5852 | -36.0573 | -41.8521
27.5913 | -36.0577 | -41.8578
27.5974 | -36.058 | -41.8636
27.6035 | -36.0584 | -41.8696
27.6096 | -36.0587 | -41.8753
27.6157 | -36.059 | -41.8807
27.6218 | -36.0594 | -41.8867
27.6279 | -36.0598 | -41.8925
27.634 | -36.0603 | -41.8985
27.6401 | -36.0606 | -41.9044
27.6462 | -36.061 | -41.9102
27.6523 | -36.0614 | -41.9161
27.6584 | -36.0618 | -41.9215
27.6645 | -36.0621 | -41.9273
27.6706 | -36.0625 | -41.9333
27.6767 | -36.0629 | -41.9393
27.6828 | -36.0633 | -41.9447
27.6889 | -36.0636 | -41.9505
27.695 | -36.064 | -41.9559
27.7011 | -36.0644 | -41.9619
27.7072 | -36.0648 | -41.9676
27.7133 | -36.0652 | -41.9733
27.7194 | -36.0655 | -41.979
27.7256 | -36.0659 | -41.9846
27.7317 | -36.0663 | -41.9903
27.7378 | -36.0667 | -41.9963
27.7439 | -36.0671 | -42.0021
27.75 | -36.0674 | -42.008
27.7561 | -36.0678 | -42.0133
27.7622 | -36.0681 | -42.0185
27.7683 | -36.0684 | -42.0243
27.7744 | -36.0689 | -42.0299
27.7805 | -36.0693 | -42.0354
27.7866 | -36.0696 | -42.0414
27.7927 | -36.07 | -42.0471
27.7988 | -36.0704 | -42.0531
27.8049 | -36.0708 | -42.0588
27.811 | -36.0711 | -42.0646
27.8171 | -36.0715 | -42.07
27.8232 | -36.0719 | -42.0751
27.8293 | -36.0723 | -42.0808
27.8354 | -36.0726 | -42.0868
27.8415 | -36.0729 | -42.0923
27.8476 | -36.0732 | -42.0979
27.8537 | -36.0736 | -42.1039
27.8598 | -36.074 | -42.1096
27.866 | -36.0744 | -42.115
27.8721 | -36.0748 | -42.1208
27.8782 | -36.0752 | -42.1265
27.8843 | -36.0755 | -42.1324
27.8904 | -36.0759 | -42.1382
27.8965 | -36.0761 | -42.1441
27.9026 | -36.0765 | -42.1498
27.9087 | -36.0769 | -42.1553
27.9148 | -36.0772 | -42.1611
27.9209 | -36.0776 | -42.1669
27.927 | -36.078 | -42.1721
27.9331 | -36.0784 | -42.1774
27.9392 | -36.0788 | -42.1832
27.9453 | -36.0791 | -42.1887
27.9514 | -36.0794 | -42.1939
27.9575 | -36.0798 | -42.1996
27.9636 | -36.0801 | -42.2052
27.9697 | -36.0804 | -42.211
27.9758 | -36.0808 | -42.2168
27.9819 | -36.0812 | -42.2228
27.988 | -36.0815 | -42.2283
27.9941 | -36.0818 | -42.234
28.0002 | -36.0822 | -42.2389
28.0063 | -36.0825 | -42.245
28.0125 | -36.0828 | -42.2509
28.0186 | -36.0831 | -42.2568
28.0247 | -36.0835 | -42.2627
28.0308 | -36.0839 | -42.2682
28.0369 | -36.0842 | -42.2733
28.043 | -36.0845 | -42.2791
28.0491 | -36.0849 | -42.2852
28.0552 | -36.0852 | -42.2909
28.0613 | -36.0856 | -42.2969
28.0674 | -36.0859 | -42.3024
28.0735 | -36.0863 | -42.3075

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.7992 | -47.6246
21.4908 | -35.9476 | -39.5051
16.4948 | -35.2155 | -31.5127
12.6602 | -34.7763 | -24.6031
9.7171 | -34.4996 | -19.107
7.4582 | -34.3109 | -14.6031
5.7244 | -34.1807 | -10.9459
4.3936 | -34.0953 | -7.9901
3.3722 | -34.0658 | -5.6566
2.5883 | -34.0533 | -3.8823
1.9866 | -34.0477 | -2.4738
1.5248 | -34.0488 | -1.3498
1.1703 | -34.0444 | -0.58861
1 | -34.0486 | -0.2237
0.89822 | -34.0443 | 0.08656
0.68941 | -34.0292 | 0.72895
0.52919 | -34.0258 | 1.2363
0.40612 | -34.0172 | 1.8792
0.31171 | -34.0115 | 2.3628
0.23928 | -34.0216 | 3.1052
0.18358 | -34.0361 | 3.5379
0.14091 | -34.0497 | 4.4214
0.10819 | -34.0672 | 5.582
0.083017 | -34.0822 | 7.0536
0.063705 | -34.1323 | 8.9142
0.048923 | -34.2032 | 11.4083
0.037479 | -34.2991 | 14.5799
0.028801 | -34.4646 | 18.6555
0.022078 | -34.7796 | 24.0787
0.017023 | -35.2757 | 30.4554
0.013018 | -36.0761 | 38.1211
0.010014 | -37.1863 | 46.0473

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003655 | -15.1275
0.00731 | -33.3662
0.010965 | -36.7892
0.014619 | -41.1719
0.018274 | -43.3426
0.021929 | -45.5429
0.025584 | -47.409
0.029239 | -48.6729
0.032894 | -50.1451
0.036548 | -51.5326
0.040203 | -52.3249
0.043858 | -52.4104
0.047513 | -53.6808
0.051168 | -54.7772
0.054823 | -55.2152
0.058477 | -55.6798
0.062132 | -57.3783
0.065787 | -56.6486
0.069442 | -57.5694
0.073097 | -58.5658
0.076752 | -60.0313
0.080406 | -60.4255
0.084061 | -60.3644
0.087716 | -61.1705
0.091371 | -61.235
0.095026 | -62.2455
0.098681 | -63.8397
0.10233 | -64.8724
0.10599 | -64.7656
0.10965 | -64.3651
0.1133 | -64.4612
0.11696 | -63.8541
0.12061 | -64.1582
0.12426 | -64.4769
0.12792 | -65.0615
0.13157 | -65.7357
0.13523 | -66.1002
0.13888 | -66.3744
0.14254 | -66.7543
0.14619 | -68.2845
0.14985 | -68.5411
0.1535 | -67.0888
0.15716 | -66.9674
0.16081 | -68.1531
0.16447 | -68.0902
0.16812 | -68.7735
0.17178 | -67.7885
0.17543 | -68.3108
0.17909 | -67.6649
0.18274 | -68.6582
0.1864 | -69.3002
0.19005 | -70.8771
0.19371 | -69.5871
0.19736 | -69.0248
0.20102 | -69.6909
0.20467 | -70.524
0.20833 | -70.9322
0.21198 | -71.3125
0.21563 | -70.2634
0.21929 | -69.8813
0.22295 | -70.1707
0.2266 | -71.3775
0.23025 | -71.4912
0.23391 | -71.736
0.23756 | -71.2733
0.24122 | -71.2901
0.24487 | -71.5338
0.24853 | -72.8405
0.25218 | -73.2513
0.25584 | -73.096
0.25949 | -73.4709
0.26315 | -73.5428
0.2668 | -74.3932
0.27046 | -74.5066
0.27411 | -73.4947
0.27777 | -73.335
0.28142 | -74.1288
0.28508 | -73.3553
0.28873 | -73.109
0.29239 | -73.9969
0.29604 | -73.547
0.2997 | -73.315
0.30335 | -74.7065
0.30701 | -74.2111
0.31066 | -73.9459
0.31432 | -74.3465
0.31797 | -75.9153
0.32163 | -75.8503
0.32528 | -75.1076
0.32893 | -73.879
0.33259 | -74.2314
0.33625 | -75.5058
0.3399 | -75.4524
0.34355 | -74.6932
0.34721 | -74.9401
0.35086 | -76.1432
0.35452 | -76.6009
0.35817 | -77.5068
0.36183 | -76.0977
0.36548 | -75.7555
0.36914 | -77.3094
0.37279 | -76.6322
0.37645 | -76.7696
0.3801 | -76.9985
0.38376 | -75.8684
0.38741 | -75.9087
0.39107 | -77.1455
0.39472 | -77.0366
0.39838 | -76.748
0.40203 | -75.754
0.40569 | -76.6276
0.40934 | -76.9927
0.413 | -76.831
0.41665 | -76.7662
0.42031 | -77.3953
0.42396 | -76.7481
0.42762 | -77.1087
0.43127 | -77.2991
0.43493 | -77.3585
0.43858 | -77.2635
0.44223 | -77.1208
0.44589 | -78.2318
0.44954 | -77.2407
0.4532 | -77.7996
0.45686 | -78.1948
0.46051 | -78.4783
0.46416 | -78.484
0.46782 | -77.6505
0.47147 | -77.6177
0.47513 | -79.5823
0.47878 | -79.3796
0.48244 | -78.665
0.48609 | -78.1303
0.48975 | -78.6441
0.4934 | -79.3575
0.49706 | -80.4863
0.50071 | -78.6531
0.50437 | -77.2948
0.50802 | -77.2687
0.51168 | -80.1254
0.51533 | -80.5319
0.51899 | -79.4562
0.52264 | -80.2044
0.5263 | -79.6784
0.52995 | -78.8724
0.53361 | -78.4711
0.53726 | -78.4412
0.54092 | -78.5936
0.54457 | -78.8599
0.54823 | -78.9847
0.55188 | -78.9557
0.55554 | -79.1705
0.55919 | -79.5867
0.56285 | -79.7728
0.5665 | -79.6665
0.57015 | -79.8521
0.57381 | -80.5627
0.57746 | -81.2984
0.58112 | -81.3696
0.58477 | -80.8853
0.58843 | -80.5163
0.59208 | -80.833
0.59574 | -81.05
0.59939 | -80.7768
0.60305 | -80.5917
0.6067 | -80.2932
0.61036 | -80.0125
0.61401 | -80.3714
0.61767 | -80.5417
0.62132 | -80.5155
0.62498 | -80.7794
0.62863 | -81.3043
0.63229 | -82.0781
0.63594 | -82.5776
0.6396 | -81.9678
0.64325 | -81.3774
0.64691 | -81.3413
0.65056 | -81.1043
0.65422 | -80.6913
0.65787 | -80.8706
0.66153 | -81.6007
0.66518 | -82.3947
0.66883 | -82.5358
0.67249 | -81.7789
0.67614 | -81.0949
0.6798 | -81.4374
0.68346 | -82.2496
0.68711 | -82.5263
0.69076 | -82.6712
0.69442 | -82.5609
0.69807 | -82.0619
0.70173 | -81.6502
0.70538 | -81.1129
0.70904 | -80.7912
0.71269 | -81.1107
0.71635 | -81.688
0.72 | -82.5206
0.72366 | -83.1975
0.72731 | -83.5555
0.73097 | -83.9921
0.73462 | -83.897
0.73828 | -82.9281
0.74193 | -82.2762
0.74559 | -82.4072
0.74924 | -82.6492
0.7529 | -82.4874
0.75655 | -82.1355
0.76021 | -82.0073
0.76386 | -82.1016
0.76752 | -82.1267
0.77117 | -82.1945
0.77482 | -81.9846
0.77848 | -81.6836
0.78214 | -82.0819
0.78579 | -83.049
0.78944 | -83.5166
0.7931 | -83.3386
0.79675 | -83.0399
0.80041 | -82.8114
0.80406 | -83.0633
0.80772 | -83.5751
0.81137 | -83.7006
0.81503 | -83.5617
0.81868 | -83.4811
0.82234 | -83.471
0.82599 | -83.4739
0.82965 | -83.2104
0.8333 | -83.1465
0.83696 | -83.2807
0.84061 | -83.5424
0.84427 | -84.0127
0.84792 | -84.2546
0.85158 | -84.1918
0.85523 | -84.2053
0.85889 | -84.5781
0.86254 | -84.5195
0.8662 | -83.6903
0.86985 | -83.1708
0.87351 | -83.2942
0.87716 | -83.37
0.88082 | -83.3781
0.88447 | -83.7055
0.88813 | -83.8175
0.89178 | -83.756
0.89543 | -84.2612
0.89909 | -84.5049
0.90275 | -84.6352
0.9064 | -85.2579
0.91005 | -85.8314
0.91371 | -85.541
0.91736 | -84.8466
0.92102 | -84.6492
0.92467 | -84.7135
0.92833 | -85.0398
0.93198 | -84.9627
0.93564 | -84.3668
0.93929 | -84.3029
0.94295 | -84.6773
0.9466 | -85.0574
0.95026 | -85.4778
0.95391 | -85.8131
0.95757 | -85.5912
0.96122 | -85.2404
0.96488 | -85.3882
0.96853 | -85.8383
0.97219 | -85.6747
0.97584 | -85.038
0.9795 | -84.6995
0.98315 | -85.1566
0.98681 | -85.8587
0.99046 | -85.7411
0.99412 | -82.7004
0.99777 | -79.5348
1.0014 | -79.0676
1.0051 | -81.5202
1.0087 | -84.7729
1.0124 | -85.9473
1.016 | -85.9369
1.0197 | -85.9074
1.0234 | -85.9966
1.027 | -86.0339
1.0307 | -86.399
1.0343 | -86.7823
1.038 | -86.4422
1.0416 | -85.9874
1.0453 | -86.2869
1.0489 | -86.4374
1.0526 | -86.2384
1.0562 | -86.2198
1.0599 | -86.2635
1.0636 | -86.3165
1.0672 | -86.2875
1.0709 | -86.0781
1.0745 | -85.8555
1.0782 | -85.8371
1.0818 | -85.9241
1.0855 | -85.8955
1.0891 | -85.7751
1.0928 | -85.7834
1.0965 | -86.0594
1.1001 | -86.5122
1.1038 | -86.9795
1.1074 | -87.3937
1.1111 | -87.5972
1.1147 | -87.4234
1.1184 | -86.96
1.122 | -86.48
1.1257 | -86.2534
1.1293 | -86.3142
1.133 | -86.4078
1.1367 | -86.3449
1.1403 | -86.276
1.144 | -86.4173
1.1476 | -86.7958
1.1513 | -87.2108
1.1549 | -87.3911
1.1586 | -87.3611
1.1622 | -87.3494
1.1659 | -87.4108
1.1695 | -87.478
1.1732 | -87.5901
1.1769 | -87.7321
1.1805 | -87.7043
1.1842 | -87.4731
1.1878 | -87.251
1.1915 | -87.2096
1.1951 | -87.2894
1.1988 | -87.303
1.2024 | -87.2159
1.2061 | -87.2187
1.2098 | -87.4546
1.2134 | -87.758
1.2171 | -87.9375
1.2207 | -88.0583
1.2244 | -88.1192
1.228 | -88.1758
1.2317 | -88.4107
1.2353 | -88.7044
1.239 | -88.7893
1.2426 | -88.6578
1.2463 | -88.3654
1.25 | -87.8899
1.2536 | -87.3759
1.2573 | -87.0727
1.2609 | -87.1379
1.2646 | -87.4728
1.2682 | -87.8071
1.2719 | -88.0331
1.2755 | -88.1687
1.2792 | -88.2681
1.2828 | -88.4895
1.2865 | -88.8933
1.2902 | -89.3454
1.2938 | -89.593
1.2975 | -89.4768
1.3011 | -89.2385
1.3048 | -89.0498
1.3084 | -88.7905
1.3121 | -88.5278
1.3157 | -88.3927
1.3194 | -88.234
1.3231 | -87.9491
1.3267 | -87.7825
1.3304 | -87.973
1.334 | -88.3614
1.3377 | -88.6346
1.3413 | -88.8026
1.345 | -88.9876
1.3486 | -89.1226
1.3523 | -89.1486
1.3559 | -89.1861
1.3596 | -89.3201
1.3633 | -89.3724
1.3669 | -89.2028
1.3706 | -88.8541
1.3742 | -88.4875
1.3779 | -88.3409
1.3815 | -88.4184
1.3852 | -88.511
1.3888 | -88.5242
1.3925 | -88.4937
1.3961 | -88.445
1.3998 | -88.446
1.4035 | -88.577
1.4071 | -88.757
1.4108 | -88.9428
1.4144 | -89.2089
1.4181 | -89.3732
1.4217 | -89.1729
1.4254 | -88.7976
1.429 | -88.7556
1.4327 | -89.3045
1.4364 | -90.0892
1.44 | -90.3986
1.4437 | -90.0708
1.4473 | -89.6293
1.451 | -89.5377
1.4546 | -89.7922
1.4583 | -90.0257
1.4619 | -90.0354
1.4656 | -89.945
1.4692 | -89.9413
1.4729 | -90.0681
1.4766 | -90.2559
1.4802 | -90.379
1.4839 | -90.316
1.4875 | -90.1382
1.4912 | -89.9779
1.4948 | -89.7629
1.4985 | -89.5015
1.5021 | -89.4111
1.5058 | -89.5111
1.5094 | -89.4787
1.5131 | -89.1909
1.5168 | -89.0064
1.5204 | -89.1654
1.5241 | -89.5444
1.5277 | -89.8054
1.5314 | -89.7275
1.535 | -89.2915
1.5387 | -88.6292
1.5423 | -88.1529
1.546 | -88.1247
1.5497 | -88.3196
1.5533 | -88.2943
1.557 | -87.8415
1.5606 | -87.4358
1.5643 | -87.6056
1.5679 | -88.3702
1.5716 | -89.3634
1.5752 | -89.995
1.5789 | -89.8702
1.5825 | -89.8992
1.5862 | -89.9746
1.5899 | -90.1261
1.5935 | -90.244
1.5972 | -90.2719
1.6008 | -90.2706
1.6045 | -90.297
1.6081 | -90.3114
1.6118 | -90.2533
1.6154 | -90.1873
1.6191 | -90.205
1.6227 | -90.2993
1.6264 | -90.3858
1.6301 | -90.4131
1.6337 | -90.4308
1.6374 | -90.4941
1.641 | -90.5891
1.6447 | -90.6426
1.6483 | -90.6083
1.652 | -90.5651
1.6556 | -90.6275
1.6593 | -90.7804
1.663 | -90.8606
1.6666 | -90.7335
1.6703 | -90.4681
1.6739 | -90.2289
1.6776 | -90.1448
1.6812 | -90.2739
1.6849 | -90.5636
1.6885 | -90.8791
1.6922 | -91.0572
1.6958 | -91.048
1.6995 | -90.9395
1.7032 | -90.8746
1.7068 | -90.9473
1.7105 | -91.0926
1.7141 | -91.1496
1.7178 | -91.0516
1.7214 | -90.8844
1.7251 | -90.7838
1.7287 | -90.8249
1.7324 | -90.9432
1.736 | -90.9994
1.7397 | -90.9345
1.7434 | -90.8508
1.747 | -90.8325
1.7507 | -90.855
1.7543 | -90.9119
1.758 | -91.077
1.7616 | -91.3593
1.7653 | -91.653
1.7689 | -91.7981
1.7726 | -91.7187
1.7763 | -91.5501
1.7799 | -91.4455
1.7836 | -91.4129
1.7872 | -91.3466
1.7909 | -91.1976
1.7945 | -91.0373
1.7982 | -90.9459
1.8018 | -90.9476
1.8055 | -91.0075
1.8091 | -91.0586
1.8128 | -91.0791
1.8165 | -91.0829
1.8201 | -91.0571
1.8238 | -90.9914
1.8274 | -90.9427
1.8311 | -90.9746
1.8347 | -91.1217
1.8384 | -91.3177
1.842 | -91.448
1.8457 | -91.4744
1.8493 | -91.4578
1.853 | -91.4647
1.8567 | -91.4729
1.8603 | -91.4455
1.864 | -91.4083
1.8676 | -91.3726
1.8713 | -91.3613
1.8749 | -91.3734
1.8786 | -91.3874
1.8822 | -91.3782
1.8859 | -91.3348
1.8896 | -91.2706
1.8932 | -91.2371
1.8969 | -91.2835
1.9005 | -91.3825
1.9042 | -91.4509
1.9078 | -91.4709
1.9115 | -91.5403
1.9151 | -91.7365
1.9188 | -91.9979
1.9224 | -92.1401
1.9261 | -92.0383
1.9298 | -91.7542
1.9334 | -91.4104
1.9371 | -91.1386
1.9407 | -91.0522
1.9444 | -91.1578
1.948 | -91.3283
1.9517 | -91.4651
1.9553 | -91.5994
1.959 | -91.7828
1.9626 | -92.0003
1.9663 | -92.2158
1.97 | -92.4169
1.9736 | -92.6078
1.9773 | -92.8016
1.9809 | -92.8654
1.9846 | -91.8542
1.9882 | -89.7805
1.9919 | -87.402
1.9955 | -85.5758
1.9992 | -84.7465
2.0029 | -85.0287
2.0065 | -86.3437
2.0102 | -88.3564
2.0138 | -90.3598
2.0175 | -91.7117
2.0211 | -92.1005
2.0248 | -92.0407
2.0284 | -91.8942
2.0321 | -91.8268
2.0357 | -91.9696
2.0394 | -92.3187
2.0431 | -92.7874
2.0467 | -93.2244
2.0504 | -93.4059
2.054 | -93.1771
2.0577 | -92.6604
2.0613 | -92.1203
2.065 | -91.7789
2.0686 | -91.7463
2.0723 | -91.9627
2.0759 | -92.2445
2.0796 | -92.4327
2.0833 | -92.474
2.0869 | -92.4245
2.0906 | -92.3445
2.0942 | -92.3259
2.0979 | -92.4166
2.1015 | -92.5643
2.1052 | -92.6417
2.1088 | -92.6365
2.1125 | -92.6845
2.1162 | -92.7585
2.1198 | -92.8493
2.1235 | -92.9065
2.1271 | -92.8951
2.1308 | -92.8105
2.1344 | -92.6857
2.1381 | -92.5612
2.1417 | -92.472
2.1454 | -92.4151
2.149 | -92.4164
2.1527 | -92.4891
2.1564 | -92.6315
2.16 | -92.8441
2.1637 | -93.101
2.1673 | -93.384
2.171 | -93.6198
2.1746 | -93.7418
2.1783 | -93.7141
2.1819 | -93.5646
2.1856 | -93.3894
2.1892 | -93.2271
2.1929 | -93.1155
2.1966 | -93.0492
2.2002 | -93.0026
2.2039 | -92.9677
2.2075 | -92.9382
2.2112 | -92.926
2.2148 | -92.9436
2.2185 | -93.0009
2.2221 | -93.0992
2.2258 | -93.2183
2.2295 | -93.3573
2.2331 | -93.4905
2.2368 | -93.5633
2.2404 | -93.5193
2.2441 | -93.3513
2.2477 | -93.1415
2.2514 | -92.9853
2.255 | -92.9213
2.2587 | -92.9407
2.2623 | -92.9917
2.266 | -93.0283
2.2697 | -93.0018
2.2733 | -92.8891
2.277 | -92.7496
2.2806 | -92.6514
2.2843 | -92.6489
2.2879 | -92.7235
2.2916 | -92.8414
2.2952 | -92.9464
2.2989 | -92.9736
2.3025 | -92.9012
2.3062 | -92.763
2.3099 | -92.6506
2.3135 | -92.6533
2.3172 | -92.7882
2.3208 | -92.9967
2.3245 | -93.1755
2.3281 | -93.2181
2.3318 | -93.0707
2.3354 | -92.8383
2.3391 | -92.6602
2.3428 | -92.6438
2.3464 | -92.7875
2.3501 | -93.0234
2.3537 | -93.2271
2.3574 | -93.3147
2.361 | -93.2846
2.3647 | -93.1821
2.3683 | -93.1003
2.372 | -93.099
2.3756 | -93.155
2.3793 | -93.2221
2.383 | -93.2386
2.3866 | -93.2209
2.3903 | -93.1957
2.3939 | -93.2048
2.3976 | -93.2745
2.4012 | -93.3968
2.4049 | -93.5329
2.4085 | -93.6302
2.4122 | -93.6856
2.4158 | -93.7252
2.4195 | -93.78
2.4232 | -93.8577
2.4268 | -93.9216
2.4305 | -93.9323
2.4341 | -93.8997
2.4378 | -93.8807
2.4414 | -93.9302
2.4451 | -94.0499
2.4487 | -94.2152
2.4524 | -94.3652
2.4561 | -94.4875
2.4597 | -94.5529
2.4634 | -94.5632
2.467 | -94.5304
2.4707 | -94.4763
2.4743 | -94.4065
2.478 | -94.2992
2.4816 | -94.1883
2.4853 | -94.0759
2.4889 | -93.9678
2.4926 | -93.9105
2.4963 | -93.9624
2.4999 | -94.1293
2.5036 | -94.3523
2.5072 | -94.572
2.5109 | -94.7495
2.5145 | -94.868
2.5182 | -94.9365
2.5218 | -94.9103
2.5255 | -94.8022
2.5291 | -94.6302
2.5328 | -94.411
2.5365 | -94.1515
2.5401 | -93.9018
2.5438 | -93.7155
2.5474 | -93.6256
2.5511 | -93.6138
2.5547 | -93.6554
2.5584 | -93.7235
2.562 | -93.8241
2.5657 | -93.9592
2.5694 | -94.0879
2.573 | -94.1406
2.5767 | -94.0927
2.5803 | -93.9525
2.584 | -93.7359
2.5876 | -93.5079
2.5913 | -93.3515
2.5949 | -93.3208
2.5986 | -93.4265
2.6022 | -93.6238
2.6059 | -93.8584
2.6096 | -94.0794
2.6132 | -94.2841
2.6169 | -94.4441
2.6205 | -94.5483
2.6242 | -94.5928
2.6278 | -94.5529
2.6315 | -94.4474
2.6351 | -94.3092
2.6388 | -94.1997
2.6424 | -94.0995
2.6461 | -94.0253
2.6498 | -93.9809
2.6534 | -93.9569
2.6571 | -93.918
2.6607 | -93.8619
2.6644 | -93.814
2.668 | -93.7997
2.6717 | -93.836
2.6753 | -93.9017
2.679 | -93.9618
2.6827 | -94.0057
2.6863 | -94.02
2.69 | -94.0123
2.6936 | -93.9775
2.6973 | -93.9277
2.7009 | -93.9072
2.7046 | -93.9379
2.7082 | -94.0289
2.7119 | -94.168
2.7155 | -94.3078
2.7192 | -94.4091
2.7229 | -94.4475
2.7265 | -94.4276
2.7302 | -94.3811
2.7338 | -94.3265
2.7375 | -94.2748
2.7411 | -94.2391
2.7448 | -94.2289
2.7484 | -94.2564
2.7521 | -94.3422
2.7557 | -94.4801
2.7594 | -94.6563
2.7631 | -94.8587
2.7667 | -95.0592
2.7704 | -95.2413
2.774 | -95.3583
2.7777 | -95.3538
2.7813 | -95.2125
2.785 | -94.9654
2.7886 | -94.6796
2.7923 | -94.4105
2.7959 | -94.2034
2.7996 | -94.0717
2.8033 | -94.0129
2.8069 | -93.9975
2.8106 | -94.0169
2.8142 | -94.0285
2.8179 | -94.0196
2.8215 | -94.0032
2.8252 | -93.9939
2.8288 | -94.0112
2.8325 | -94.0416
2.8362 | -94.0669
2.8398 | -94.0748
2.8435 | -94.0444
2.8471 | -94
2.8508 | -93.9629
2.8544 | -93.9457
2.8581 | -93.9499
2.8617 | -93.9756
2.8654 | -93.9926
2.869 | -93.9969
2.8727 | -93.9926
2.8764 | -93.9955
2.88 | -94.0306
2.8837 | -94.0894
2.8873 | -94.1631
2.891 | -94.2477
2.8946 | -94.3287
2.8983 | -94.4252
2.9019 | -94.5512
2.9056 | -94.6871
2.9093 | -94.8249
2.9129 | -94.9654
2.9166 | -95.0858
2.9202 | -95.1856
2.9239 | -95.2513
2.9275 | -95.2973
2.9312 | -95.3524
2.9348 | -95.4227
2.9385 | -95.5003
2.9421 | -95.5677
2.9458 | -95.5969
2.9495 | -95.5701
2.9531 | -95.4847
2.9568 | -95.3488
2.9604 | -95.1865
2.9641 | -95.0527
2.9677 | -94.9518
2.9714 | -94.8931
2.975 | -94.8299
2.9787 | -94.7057
2.9823 | -94.5392
2.986 | -94.3227
2.9897 | -94.0773
2.9933 | -93.8557
2.997 | -93.7284
3.0006 | -93.7327
3.0043 | -93.8641
3.0079 | -94.0772
3.0116 | -94.3038
3.0152 | -94.4917
3.0189 | -94.6138
3.0225 | -94.6759
3.0262 | -94.702
3.0299 | -94.6753
3.0335 | -94.6477
3.0372 | -94.6525
3.0408 | -94.6872
3.0445 | -94.7476
3.0481 | -94.8351
3.0518 | -94.9045
3.0554 | -94.9251
3.0591 | -94.8622
3.0628 | -94.7094
3.0664 | -94.445
3.0701 | -94.1372
3.0737 | -93.8515
3.0774 | -93.6602
3.081 | -93.6116
3.0847 | -93.7033
3.0883 | -93.9116
3.092 | -94.1928
3.0956 | -94.4591
3.0993 | -94.6428
3.103 | -94.6965
3.1066 | -94.595
3.1103 | -94.3708
3.1139 | -94.1019
3.1176 | -93.8305
3.1212 | -93.6268
3.1249 | -93.5247
3.1285 | -93.5443
3.1322 | -93.6902
3.1359 | -93.9384
3.1395 | -94.2495
3.1432 | -94.601
3.1468 | -94.9422
3.1505 | -95.2246
3.1541 | -95.4329
3.1578 | -95.54
3.1614 | -95.4683
3.1651 | -95.4248
3.1687 | -95.3375
3.1724 | -95.2197
3.1761 | -95.1106
3.1797 | -95.0221
3.1834 | -94.9683
3.187 | -94.9585
3.1907 | -94.9775
3.1943 | -95.0107
3.198 | -95.0494
3.2016 | -95.0877
3.2053 | -95.1388
3.2089 | -95.2033
3.2126 | -95.2693
3.2163 | -95.3375
3.2199 | -95.3846
3.2236 | -95.4208
3.2272 | -95.4648
3.2309 | -95.5164
3.2345 | -95.5867
3.2382 | -95.6741
3.2418 | -95.766
3.2455 | -95.8464
3.2492 | -95.9121
3.2528 | -95.9568
3.2565 | -95.9623
3.2601 | -95.9243
3.2638 | -95.8656
3.2674 | -95.7934
3.2711 | -95.7272
3.2747 | -95.6745
3.2784 | -95.638
3.282 | -95.6284
3.2857 | -95.6659
3.2894 | -95.7367
3.293 | -95.823
3.2967 | -95.9105
3.3003 | -95.9805
3.304 | -96.0082
3.3076 | -96.0071
3.3113 | -95.9907
3.3149 | -95.9806
3.3186 | -95.9911
3.3222 | -96.0387
3.3259 | -96.1159
3.3296 | -96.2283
3.3332 | -96.3403
3.3369 | -96.4289
3.3405 | -96.4613
3.3442 | -96.4444
3.3478 | -96.3946
3.3515 | -96.3217
3.3551 | -96.2698
3.3588 | -96.2542
3.3624 | -96.2711
3.3661 | -96.3339
3.3698 | -96.4219
3.3734 | -96.5057
3.3771 | -96.573
3.3807 | -96.593
3.3844 | -96.5543
3.388 | -96.4765
3.3917 | -96.3648
3.3953 | -96.2559
3.399 | -96.172
3.4027 | -96.1238
3.4063 | -96.1244
3.41 | -96.1671
3.4136 | -96.2399
3.4173 | -96.3064
3.4209 | -96.3266
3.4246 | -96.2972
3.4282 | -96.2203
3.4319 | -96.1193
3.4355 | -96.0336
3.4392 | -95.9808
3.4429 | -95.9762
3.4465 | -96.0261
3.4502 | -96.1161
3.4538 | -96.215
3.4575 | -96.3025
3.4611 | -96.3594
3.4648 | -96.3604
3.4684 | -96.3104
3.4721 | -96.2388
3.4758 | -96.1598
3.4794 | -96.0881
3.4831 | -96.0437
3.4867 | -96.0349
3.4904 | -96.0593
3.494 | -96.1026
3.4977 | -96.146
3.5013 | -96.1823
3.505 | -96.2071
3.5086 | -96.2389
3.5123 | -96.279
3.516 | -96.3349
3.5196 | -96.4127
3.5233 | -96.4918
3.5269 | -96.5569
3.5306 | -96.5823
3.5342 | -96.5656
3.5379 | -96.5162
3.5415 | -96.4416
3.5452 | -96.3786
3.5488 | -96.3473
3.5525 | -96.3502
3.5562 | -96.3913
3.5598 | -96.477
3.5635 | -96.596
3.5671 | -96.7322
3.5708 | -96.8628
3.5744 | -96.9813
3.5781 | -97.0773
3.5817 | -97.1253
3.5854 | -97.141
3.589 | -97.1205
3.5927 | -97.0781
3.5964 | -97.0332
3.6 | -96.9984
3.6037 | -96.9874
3.6073 | -96.9982
3.611 | -97.0215
3.6146 | -97.0494
3.6183 | -97.0855
3.6219 | -97.1421
3.6256 | -97.2125
3.6293 | -97.2661
3.6329 | -97.2978
3.6366 | -97.3001
3.6402 | -97.2802
3.6439 | -97.2414
3.6475 | -97.2084
3.6512 | -97.1927
3.6548 | -97.1941
3.6585 | -97.2178
3.6621 | -97.2444
3.6658 | -97.2473
3.6695 | -97.2078
3.6731 | -97.1296
3.6768 | -97.047
3.6804 | -96.9886
3.6841 | -96.9753
3.6877 | -97.002
3.6914 | -97.0074
3.695 | -97.0103
3.6987 | -96.9899
3.7024 | -96.9651
3.706 | -96.9315
3.7097 | -96.9057
3.7133 | -96.8982
3.717 | -96.9125
3.7206 | -96.9605
3.7243 | -97.0348
3.7279 | -97.1243
3.7316 | -97.2009
3.7352 | -97.2293
3.7389 | -97.2095
3.7426 | -97.1408
3.7462 | -97.0344
3.7499 | -96.9011
3.7535 | -96.7505
3.7572 | -96.6051
3.7608 | -96.4744
3.7645 | -96.3457
3.7681 | -96.2116
3.7718 | -96.0655
3.7754 | -95.9129
3.7791 | -95.7501
3.7828 | -95.5888
3.7864 | -95.4518
3.7901 | -95.3595
3.7937 | -95.3345
3.7974 | -95.3926
3.801 | -95.5343
3.8047 | -95.7483
3.8083 | -96.0037
3.812 | -96.2489
3.8157 | -96.4314
3.8193 | -96.509
3.823 | -96.4581
3.8266 | -96.3096
3.8303 | -96.092
3.8339 | -95.851
3.8376 | -95.6354
3.8412 | -95.4943
3.8449 | -95.4416
3.8485 | -95.4832
3.8522 | -95.6041
3.8559 | -95.7813
3.8595 | -95.9828
3.8632 | -96.1822
3.8668 | -96.3434
3.8705 | -96.4295
3.8741 | -96.4397
3.8778 | -96.3847
3.8814 | -96.2658
3.8851 | -96.1096
3.8887 | -95.9347
3.8924 | -95.7782
3.8961 | -95.669
3.8997 | -95.6196
3.9034 | -95.6435
3.907 | -95.7492
3.9107 | -95.9257
3.9143 | -96.1477
3.918 | -96.3845
3.9216 | -96.6097
3.9253 | -96.8071
3.9289 | -96.9833
3.9326 | -97.1546
3.9363 | -97.3067
3.9399 | -97.4498
3.9436 | -97.5939
3.9472 | -97.7297
3.9509 | -97.8421
3.9545 | -97.9339
3.9582 | -97.9919
3.9618 | -98.0225
3.9655 | -98.0191
3.9692 | -97.8105
3.9728 | -97.4537
3.9765 | -97.0406
3.9801 | -96.5224
3.9838 | -95.9472
3.9874 | -95.3814
3.9911 | -94.8975
3.9947 | -94.5442
3.9984 | -94.3619
4.002 | -94.3641
4.0057 | -94.5475
4.0094 | -94.8858
4.013 | -95.3415
4.0167 | -95.8443
4.0203 | -96.3296
4.024 | -96.7441
4.0276 | -97.0695
4.0313 | -97.3796
4.0349 | -97.5724
4.0386 | -97.6364
4.0423 | -97.6929
4.0459 | -97.7402
4.0496 | -97.7837
4.0532 | -97.8275
4.0569 | -97.8737
4.0605 | -97.9088
4.0642 | -97.9318
4.0678 | -97.9454
4.0715 | -97.946
4.0751 | -97.9236
4.0788 | -97.8895
4.0825 | -97.8441
4.0861 | -97.7924
4.0898 | -97.7563
4.0934 | -97.7525
4.0971 | -97.7859
4.1007 | -97.8575
4.1044 | -97.9575
4.108 | -98.0762
4.1117 | -98.1904
4.1153 | -98.2926
4.119 | -98.3807
4.1227 | -98.4429
4.1263 | -98.4735
4.13 | -98.4719
4.1336 | -98.4472
4.1373 | -98.4156
4.1409 | -98.3987
4.1446 | -98.3959
4.1482 | -98.3987
4.1519 | -98.3841
4.1555 | -98.3528
4.1592 | -98.302
4.1629 | -98.2331
4.1665 | -98.1508
4.1702 | -98.0561
4.1738 | -97.9662
4.1775 | -97.8812
4.1811 | -97.8005
4.1848 | -97.7296
4.1884 | -97.6594
4.1921 | -97.6073
4.1958 | -97.571
4.1994 | -97.5549
4.2031 | -97.5572
4.2067 | -97.5753
4.2104 | -97.5982
4.214 | -97.6144
4.2177 | -97.6259
4.2213 | -97.6455
4.225 | -97.6835
4.2286 | -97.7308
4.2323 | -97.7989
4.236 | -97.8812
4.2396 | -97.9544
4.2433 | -98.0088
4.2469 | -98.0381
4.2506 | -98.0367
4.2542 | -98.0114
4.2579 | -97.9592
4.2615 | -97.8981
4.2652 | -97.8451
4.2688 | -97.8114
4.2725 | -97.7941
4.2762 | -97.7876
4.2798 | -97.7922
4.2835 | -97.8026
4.2871 | -97.8158
4.2908 | -97.8203
4.2944 | -97.8231
4.2981 | -97.8431
4.3017 | -97.8719
4.3054 | -97.9198
4.3091 | -97.9908
4.3127 | -98.076
4.3164 | -98.1771
4.32 | -98.2899
4.3237 | -98.4051
4.3273 | -98.5151
4.331 | -98.6067
4.3346 | -98.6649
4.3383 | -98.683
4.3419 | -98.662
4.3456 | -98.6091
4.3493 | -98.5279
4.3529 | -98.4202
4.3566 | -98.3094
4.3602 | -98.2074
4.3639 | -98.1105
4.3675 | -98.029
4.3712 | -97.9842
4.3748 | -97.966
4.3785 | -97.9702
4.3822 | -97.9885
4.3858 | -98.0122
4.3895 | -98.0425
4.3931 | -98.0844
4.3968 | -98.131
4.4004 | -98.1883
4.4041 | -98.2534
4.4077 | -98.3253
4.4114 | -98.4015
4.415 | -98.4745
4.4187 | -98.5457
4.4224 | -98.6171
4.426 | -98.6848
4.4297 | -98.7557
4.4333 | -98.8309
4.437 | -98.9072
4.4406 | -98.9706
4.4443 | -99.0366
4.4479 | -99.1096
4.4516 | -99.1847
4.4552 | -99.2378
4.4589 | -99.2748
4.4626 | -99.28
4.4662 | -99.2547
4.4699 | -99.1884
4.4735 | -99.0917
4.4772 | -98.976
4.4808 | -98.8634
4.4845 | -98.7616
4.4881 | -98.6744
4.4918 | -98.6189
4.4954 | -98.5709
4.4991 | -98.5313
4.5028 | -98.506
4.5064 | -98.4851
4.5101 | -98.4693
4.5137 | -98.4494
4.5174 | -98.4193
4.521 | -98.3811
4.5247 | -98.3505
4.5283 | -98.3205
4.532 | -98.3058
4.5357 | -98.3072
4.5393 | -98.3107
4.543 | -98.3309
4.5466 | -98.3644
4.5503 | -98.3884
4.5539 | -98.4074
4.5576 | -98.4129
4.5612 | -98.402
4.5649 | -98.39
4.5685 | -98.3633
4.5722 | -98.3239
4.5759 | -98.2761
4.5795 | -98.2211
4.5832 | -98.1587
4.5868 | -98.093
4.5905 | -98.0256
4.5941 | -97.949
4.5978 | -97.8717
4.6014 | -97.8004
4.6051 | -97.7482
4.6087 | -97.7353
4.6124 | -97.7508
4.6161 | -97.8122
4.6197 | -97.9262
4.6234 | -98.0756
4.627 | -98.2523
4.6307 | -98.4298
4.6343 | -98.5933
4.638 | -98.7286
4.6416 | -98.8216
4.6453 | -98.8456
4.649 | -98.7982
4.6526 | -98.6911
4.6563 | -98.5323
4.6599 | -98.3284
4.6636 | -98.0971
4.6672 | -97.8718
4.6709 | -97.6778
4.6745 | -97.5249
4.6782 | -97.4203
4.6818 | -97.3763
4.6855 | -97.3905
4.6892 | -97.4703
4.6928 | -97.6065
4.6965 | -97.791
4.7001 | -98.0145
4.7038 | -98.2514
4.7074 | -98.4826
4.7111 | -98.6857
4.7147 | -98.8649
4.7184 | -99.0152
4.7221 | -99.1199
4.7257 | -99.1982
4.7294 | -99.275
4.733 | -99.355
4.7367 | -99.4186
4.7403 | -99.4681
4.744 | -99.5008
4.7476 | -99.5287
4.7513 | -99.5511
4.7549 | -99.5641
4.7586 | -99.5642
4.7623 | -99.5407
4.7659 | -99.5123
4.7696 | -99.4742
4.7732 | -99.4275
4.7769 | -99.3887
4.7805 | -99.3609
4.7842 | -99.3477
4.7878 | -99.3425
4.7915 | -99.3469
4.7951 | -99.3677
4.7988 | -99.3937
4.8025 | -99.4246
4.8061 | -99.4615
4.8098 | -99.4886
4.8134 | -99.4932
4.8171 | -99.4721
4.8207 | -99.4221
4.8244 | -99.3508
4.828 | -99.2632
4.8317 | -99.1741
4.8354 | -99.0952
4.839 | -99.0343
4.8427 | -98.9803
4.8463 | -98.9426
4.85 | -98.929
4.8536 | -98.9397
4.8573 | -98.9773
4.8609 | -99.0382
4.8646 | -99.1273
4.8682 | -99.2376
4.8719 | -99.3631
4.8756 | -99.4961
4.8792 | -99.6192
4.8829 | -99.7262
4.8865 | -99.8192
4.8902 | -99.8812
4.8938 | -99.8954
4.8975 | -99.8694
4.9011 | -99.8076
4.9048 | -99.7209
4.9084 | -99.6145
4.9121 | -99.499
4.9158 | -99.38
4.9194 | -99.2738
4.9231 | -99.1802
4.9267 | -99.1034
4.9304 | -99.0363
4.934 | -98.9898
4.9377 | -98.969
4.9413 | -98.9622
4.945 | -98.9704
4.9486 | -98.999
4.9523 | -99.0391
4.956 | -99.0827
4.9596 | -99.1125
4.9633 | -99.0263
4.9669 | -98.9346
4.9706 | -98.8058
4.9742 | -98.611
4.9779 | -98.354
4.9815 | -98.0536
4.9852 | -97.7504
4.9889 | -97.4791
4.9925 | -97.2656
4.9962 | -97.1326
4.9998 | -97.0947
5.0035 | -97.1545
5.0071 | -97.3133
5.0108 | -97.5504
5.0144 | -97.837
5.0181 | -98.1523
5.0217 | -98.4646
5.0254 | -98.7446
5.0291 | -98.9804
5.0327 | -99.1676
5.0364 | -99.3039
5.04 | -99.4478
5.0437 | -99.4838
5.0473 | -99.4936
5.051 | -99.5112
5.0546 | -99.5398
5.0583 | -99.5821
5.0619 | -99.6251
5.0656 | -99.6713
5.0693 | -99.7257
5.0729 | -99.7771
5.0766 | -99.8185
5.0802 | -99.8458
5.0839 | -99.8494
5.0875 | -99.8287
5.0912 | -99.7815
5.0948 | -99.7014
5.0985 | -99.5962
5.1022 | -99.4729
5.1058 | -99.3444
5.1095 | -99.232
5.1131 | -99.1375
5.1168 | -99.073
5.1204 | -99.0354
5.1241 | -99.0301
5.1277 | -99.062
5.1314 | -99.1172
5.135 | -99.1858
5.1387 | -99.2642
5.1424 | -99.3464
5.146 | -99.42
5.1497 | -99.4712
5.1533 | -99.5
5.157 | -99.4973
5.1606 | -99.4646
5.1643 | -99.4053
5.1679 | -99.335
5.1716 | -99.2567
5.1752 | -99.1745
5.1789 | -99.0989
5.1826 | -99.033
5.1862 | -98.9745
5.1899 | -98.9368
5.1935 | -98.9157
5.1972 | -98.9165
5.2008 | -98.9498
5.2045 | -99.0233
5.2081 | -99.14
5.2118 | -99.2944
5.2155 | -99.4728
5.2191 | -99.6589
5.2228 | -99.8462
5.2264 | -100.0191
5.2301 | -100.1555
5.2337 | -100.2486
5.2374 | -100.2929
5.241 | -100.2916
5.2447 | -100.2452
5.2483 | -100.1688
5.252 | -100.0799
5.2557 | -99.9867
5.2593 | -99.8975
5.263 | -99.8183
5.2666 | -99.7474
5.2703 | -99.6974
5.2739 | -99.6313
5.2776 | -99.5739
5.2812 | -99.5227
5.2849 | -99.4629
5.2885 | -99.4049
5.2922 | -99.359
5.2959 | -99.3203
5.2995 | -99.2875
5.3032 | -99.2644
5.3068 | -99.247
5.3105 | -99.2339
5.3141 | -99.2285
5.3178 | -99.2347
5.3214 | -99.2505
5.3251 | -99.2762
5.3288 | -99.3119
5.3324 | -99.3516
5.3361 | -99.3914
5.3397 | -99.4227
5.3434 | -99.4163
5.347 | -99.395
5.3507 | -99.3428
5.3543 | -99.2614
5.358 | -99.1604
5.3616 | -99.0306
5.3653 | -98.8892
5.369 | -98.7596
5.3726 | -98.6438
5.3763 | -98.5467
5.3799 | -98.4837
5.3836 | -98.4658
5.3872 | -98.497
5.3909 | -98.5771
5.3945 | -98.6929
5.3982 | -98.8343
5.4018 | -98.9936
5.4055 | -99.1538
5.4092 | -99.303
5.4128 | -99.4202
5.4165 | -99.5
5.4201 | -99.5343
5.4238 | -99.5064
5.4274 | -99.439
5.4311 | -99.3169
5.4347 | -99.1512
5.4384 | -98.9624
5.4421 | -98.7513
5.4457 | -98.5337
5.4494 | -98.324
5.453 | -98.1342
5.4567 | -97.9807
5.4603 | -97.8685
5.464 | -97.8049
5.4676 | -97.7912
5.4713 | -97.8264
5.4749 | -97.9149
5.4786 | -98.0551
5.4823 | -98.2382
5.4859 | -98.4532
5.4896 | -98.6962
5.4932 | -98.9547
5.4969 | -99.2203
5.5005 | -99.4731
5.5042 | -99.6991
5.5078 | -99.8998
5.5115 | -100.0566
5.5151 | -100.16
5.5188 | -100.2165
5.5225 | -100.2257
5.5261 | -100.2139
5.5298 | -100.1808
5.5334 | -100.1228
5.5371 | -100.0519
5.5407 | -99.9685
5.5444 | -99.8701
5.548 | -99.7718
5.5517 | -99.6784
5.5554 | -99.59
5.559 | -99.5007
5.5627 | -99.4241
5.5663 | -99.3605
5.57 | -99.3066
5.5736 | -99.2607
5.5773 | -99.2222
5.5809 | -99.1982
5.5846 | -99.1875
5.5882 | -99.1886
5.5919 | -99.2041
5.5956 | -99.234
5.5992 | -99.2718
5.6029 | -99.3236
5.6065 | -99.3776
5.6102 | -99.4233
5.6138 | -99.4574
5.6175 | -99.4868
5.6211 | -99.5021
5.6248 | -99.5059
5.6284 | -99.5131
5.6321 | -99.5059
5.6358 | -99.4944
5.6394 | -99.4862
5.6431 | -99.4833
5.6467 | -99.4779
5.6504 | -99.4692
5.654 | -99.4614
5.6577 | -99.4496
5.6613 | -99.4337
5.665 | -99.4078
5.6687 | -99.3771
5.6723 | -99.3415
5.676 | -99.2927
5.6796 | -99.2343
5.6833 | -99.1741
5.6869 | -99.1191
5.6906 | -99.0736
5.6942 | -99.0429
5.6979 | -99.0283
5.7015 | -99.0197
5.7052 | -99.0219
5.7089 | -99.0377
5.7125 | -99.0555
5.7162 | -99.076
5.7198 | -99.0955
5.7235 | -99.1099
5.7271 | -99.1233
5.7308 | -99.1283
5.7344 | -99.1258
5.7381 | -99.1084
5.7417 | -99.0744
5.7454 | -99.0147
5.7491 | -98.9323
5.7527 | -98.8255
5.7564 | -98.7014
5.76 | -98.5789
5.7637 | -98.4582
5.7673 | -98.344
5.771 | -98.241
5.7746 | -98.1523
5.7783 | -98.0823
5.782 | -98.0382
5.7856 | -98.0192
5.7893 | -98.0247
5.7929 | -98.0451
5.7966 | -98.1037
5.8002 | -98.1414
5.8039 | -98.1873
5.8075 | -98.2287
5.8112 | -98.2602
5.8148 | -98.2825
5.8185 | -98.2891
5.8222 | -98.2743
5.8258 | -98.2449
5.8295 | -98.2038
5.8331 | -98.1605
5.8368 | -98.1197
5.8404 | -98.0855
5.8441 | -98.0585
5.8477 | -98.0289
5.8514 | -98.0028
5.855 | -97.9808
5.8587 | -97.9631
5.8624 | -97.9473
5.866 | -97.9295
5.8697 | -97.9022
5.8733 | -97.871
5.877 | -97.832
5.8806 | -97.7889
5.8843 | -97.7417
5.8879 | -97.6926
5.8916 | -97.6497
5.8953 | -97.6145
5.8989 | -97.5865
5.9026 | -97.5729
5.9062 | -97.5679
5.9099 | -97.5774
5.9135 | -97.6042
5.9172 | -97.643
5.9208 | -97.6913
5.9245 | -97.7453
5.9281 | -97.8084
5.9318 | -97.8807
5.9355 | -97.9567
5.9391 | -98.0346
5.9428 | -98.1091
5.9464 | -98.1847
5.9501 | -98.2507
5.9537 | -98.2995
5.9574 | -98.3311
5.961 | -98.3678
5.9647 | -98.3944
5.9683 | -98.4082
5.972 | -98.4092
5.9757 | -98.3995
5.9793 | -98.3795
5.983 | -98.3562
5.9866 | -98.3373
5.9903 | -98.3359
5.9939 | -98.3498
5.9976 | -98.3844
6.0012 | -98.442
6.0049 | -98.523
6.0086 | -98.6236
6.0122 | -98.7398
6.0159 | -98.8621
6.0195 | -98.9841
6.0232 | -99.1025
6.0268 | -99.2127
6.0305 | -99.3091
6.0341 | -99.3935
6.0378 | -99.45
6.0414 | -99.484
6.0451 | -99.5121
6.0488 | -99.529
6.0524 | -99.5172
6.0561 | -99.499
6.0597 | -99.4815
6.0634 | -99.4693
6.067 | -99.4602
6.0707 | -99.4587
6.0743 | -99.4656
6.078 | -99.4868
6.0816 | -99.52
6.0853 | -99.5631
6.089 | -99.6114
6.0926 | -99.6686
6.0963 | -99.7283
6.0999 | -99.7835
6.1036 | -99.8286
6.1072 | -99.8582
6.1109 | -99.8622
6.1145 | -99.8539
6.1182 | -99.8255
6.1219 | -99.779
6.1255 | -99.7032
6.1292 | -99.6155
6.1328 | -99.5215
6.1365 | -99.418
6.1401 | -99.3119
6.1438 | -99.2103
6.1474 | -99.1284
6.1511 | -99.0672
6.1547 | -99.0291
6.1584 | -99.0195
6.1621 | -99.031
6.1657 | -99.0645
6.1694 | -99.1114
6.173 | -99.1688
6.1767 | -99.229
6.1803 | -99.2821
6.184 | -99.3271
6.1876 | -99.3613
6.1913 | -99.3876
6.1949 | -99.4024
6.1986 | -99.4124
6.2023 | -99.4092
6.2059 | -99.3856
6.2096 | -99.3701
6.2132 | -99.3524
6.2169 | -99.337
6.2205 | -99.3169
6.2242 | -99.3007
6.2278 | -99.2877
6.2315 | -99.2826
6.2352 | -99.2859
6.2388 | -99.2984
6.2425 | -99.3212
6.2461 | -99.3523
6.2498 | -99.3957
6.2534 | -99.4607
6.2571 | -99.5359
6.2607 | -99.6276
6.2644 | -99.7375
6.268 | -99.8596
6.2717 | -99.9936
6.2754 | -100.1237
6.279 | -100.2499
6.2827 | -100.3685
6.2863 | -100.4773
6.29 | -100.5625
6.2936 | -100.637
6.2973 | -100.6963
6.3009 | -100.7328
6.3046 | -100.7613
6.3082 | -100.7947
6.3119 | -100.8248
6.3156 | -100.8457
6.3192 | -100.8698
6.3229 | -100.897
6.3265 | -100.9467
6.3302 | -100.9869
6.3338 | -101.0219
6.3375 | -101.0634
6.3411 | -101.1187
6.3448 | -101.1837
6.3485 | -101.2396
6.3521 | -101.2942
6.3558 | -101.3463
6.3594 | -101.3896
6.3631 | -101.4277
6.3667 | -101.4505
6.3704 | -101.4626
6.374 | -101.4581
6.3777 | -101.435
6.3813 | -101.409
6.385 | -101.3889
6.3887 | -101.3626
6.3923 | -101.3384
6.396 | -101.3096
6.3996 | -101.2833
6.4033 | -101.2498
6.4069 | -101.2233
6.4106 | -101.2049
6.4142 | -101.1877
6.4179 | -101.1749
6.4215 | -101.1717
6.4252 | -101.1744
6.4289 | -101.1786
6.4325 | -101.1953
6.4362 | -101.2133
6.4398 | -101.2293
6.4435 | -101.2406
6.4471 | -101.257
6.4508 | -101.2751
6.4544 | -101.2837
6.4581 | -101.2834
6.4618 | -101.2836
6.4654 | -101.2806
6.4691 | -101.2742
6.4727 | -101.2725
6.4764 | -101.2776
6.48 | -101.28
6.4837 | -101.2793
6.4873 | -101.2864
6.491 | -101.2995
6.4946 | -101.3216
6.4983 | -101.3514
6.502 | -101.3863
6.5056 | -101.4206
6.5093 | -101.4528
6.5129 | -101.4835
6.5166 | -101.5133
6.5202 | -101.5374
6.5239 | -101.5559
6.5275 | -101.5726
6.5312 | -101.587
6.5348 | -101.5947
6.5385 | -101.5965
6.5422 | -101.6021
6.5458 | -101.6076
6.5495 | -101.6028
6.5531 | -101.5911
6.5568 | -101.5809
6.5604 | -101.567
6.5641 | -101.5513
6.5677 | -101.5305
6.5714 | -101.5054
6.5751 | -101.4771
6.5787 | -101.444
6.5824 | -101.4189
6.586 | -101.3961
6.5897 | -101.3724
6.5933 | -101.3504
6.597 | -101.3302
6.6006 | -101.3166
6.6043 | -101.3059
6.6079 | -101.2918
6.6116 | -101.275
6.6153 | -101.259
6.6189 | -101.2433
6.6226 | -101.2242
6.6262 | -101.2074
6.6299 | -101.1873
6.6335 | -101.178
6.6372 | -101.1772
6.6408 | -101.1769
6.6445 | -101.1753
6.6481 | -101.1751
6.6518 | -101.1755
6.6555 | -101.1762
6.6591 | -101.1727
6.6628 | -101.1659
6.6664 | -101.1609
6.6701 | -101.1608
6.6737 | -101.1635
6.6774 | -101.1765
6.681 | -101.1866
6.6847 | -101.195
6.6884 | -101.212
6.692 | -101.2301
6.6957 | -101.2504
6.6993 | -101.2671
6.703 | -101.285
6.7066 | -101.3071
6.7103 | -101.3341
6.7139 | -101.3574
6.7176 | -101.3807
6.7212 | -101.4097
6.7249 | -101.4308
6.7286 | -101.4491
6.7322 | -101.4618
6.7359 | -101.471
6.7395 | -101.4843
6.7432 | -101.4981
6.7468 | -101.509
6.7505 | -101.5134
6.7541 | -101.5144
6.7578 | -101.5202
6.7614 | -101.5355
6.7651 | -101.5449
6.7688 | -101.5561
6.7724 | -101.5723
6.7761 | -101.5921
6.7797 | -101.6169
6.7834 | -101.6414
6.787 | -101.6497
6.7907 | -101.6539
6.7943 | -101.6535
6.798 | -101.6433
6.8017 | -101.6276
6.8053 | -101.6167
6.809 | -101.6014
6.8126 | -101.583
6.8163 | -101.5714
6.8199 | -101.5566
6.8236 | -101.548
6.8272 | -101.5382
6.8309 | -101.5207
6.8345 | -101.5011
6.8382 | -101.4781
6.8419 | -101.4455
6.8455 | -101.4108
6.8492 | -101.3789
6.8528 | -101.3229
6.8565 | -101.2739
6.8601 | -101.2246
6.8638 | -101.1597
6.8674 | -101.0918
6.8711 | -101.019
6.8747 | -100.9328
6.8784 | -100.8396
6.8821 | -100.7403
6.8857 | -100.642
6.8894 | -100.5416
6.893 | -100.4372
6.8967 | -100.3352
6.9003 | -100.2397
6.904 | -100.1565
6.9076 | -100.0878
6.9113 | -100.0398
6.915 | -100.0189
6.9186 | -100.0233
6.9223 | -100.0525
6.9259 | -100.1041
6.9296 | -100.1788
6.9332 | -100.2777
6.9369 | -100.3916
6.9405 | -100.5173
6.9442 | -100.6531
6.9478 | -100.7723
6.9515 | -100.8774
6.9552 | -100.9771
6.9588 | -101.0585
6.9625 | -101.1106
6.9661 | -101.1251
6.9698 | -101.1175
6.9734 | -101.0847
6.9771 | -101.0119
6.9807 | -100.921
6.9844 | -100.8149
6.988 | -100.6962
6.9917 | -100.5699
6.9954 | -100.4492
6.999 | -100.3383
7.0027 | -100.2357
7.0063 | -100.1422
7.01 | -100.0583
7.0136 | -99.9909
7.0173 | -99.9457
7.0209 | -99.9223
7.0246 | -99.9168
7.0283 | -99.9305
7.0319 | -99.9672
7.0356 | -100.0249
7.0392 | -100.0976
7.0429 | -100.1833
7.0465 | -100.279
7.0502 | -100.3841
7.0538 | -100.5018
7.0575 | -100.6181
7.0611 | -100.7269
7.0648 | -100.8375
7.0685 | -100.9477
7.0721 | -101.0643
7.0758 | -101.1813
7.0794 | -101.2886
7.0831 | -101.401
7.0867 | -101.5038
7.0904 | -101.5924
7.094 | -101.6711
7.0977 | -101.7386
7.1013 | -101.7958
7.105 | -101.8345
7.1087 | -101.8615
7.1123 | -101.8858
7.116 | -101.8983
7.1196 | -101.8963
7.1233 | -101.8901
7.1269 | -101.8769
7.1306 | -101.8619
7.1342 | -101.8517
7.1379 | -101.8321
7.1416 | -101.815
7.1452 | -101.8029
7.1489 | -101.7912
7.1525 | -101.779
7.1562 | -101.7714
7.1598 | -101.7687
7.1635 | -101.7683
7.1671 | -101.7702
7.1708 | -101.7717
7.1744 | -101.7755
7.1781 | -101.7792
7.1818 | -101.7886
7.1854 | -101.8059
7.1891 | -101.8291
7.1927 | -101.8584
7.1964 | -101.8896
7.2 | -101.9283
7.2037 | -101.9705
7.2073 | -102.0113
7.211 | -102.0557
7.2146 | -102.1018
7.2183 | -102.1458
7.222 | -102.1929
7.2256 | -102.2353
7.2293 | -102.2697
7.2329 | -102.3031
7.2366 | -102.3274
7.2402 | -102.3393
7.2439 | -102.3425
7.2475 | -102.3343
7.2512 | -102.3175
7.2549 | -102.2994
7.2585 | -102.2728
7.2622 | -102.2423
7.2658 | -102.2047
7.2695 | -102.1658
7.2731 | -102.1313
7.2768 | -102.1016
7.2804 | -102.0743
7.2841 | -102.0517
7.2877 | -102.0299
7.2914 | -102.0112
7.2951 | -101.9954
7.2987 | -101.9867
7.3024 | -101.9844
7.306 | -101.9828
7.3097 | -101.9885
7.3133 | -101.9999
7.317 | -102.0106
7.3206 | -102.018
7.3243 | -102.0258
7.3279 | -102.0272
7.3316 | -102.0347
7.3353 | -102.0425
7.3389 | -102.0445
7.3426 | -102.0489
7.3462 | -102.048
7.3499 | -102.0453
7.3535 | -102.0428
7.3572 | -102.0419
7.3608 | -102.0357
7.3645 | -102.0361
7.3682 | -102.0403
7.3718 | -102.0543
7.3755 | -102.0719
7.3791 | -102.1105
7.3828 | -102.1341
7.3864 | -102.16
7.3901 | -102.1928
7.3937 | -102.2286
7.3974 | -102.2713
7.401 | -102.3169
7.4047 | -102.355
7.4084 | -102.3838
7.412 | -102.4138
7.4157 | -102.4401
7.4193 | -102.445
7.423 | -102.439
7.4266 | -102.4209
7.4303 | -102.3917
7.4339 | -102.349
7.4376 | -102.2978
7.4412 | -102.2387
7.4449 | -102.175
7.4486 | -102.1152
7.4522 | -102.0624
7.4559 | -102.0223
7.4595 | -101.9879
7.4632 | -101.9663
7.4668 | -101.9607
7.4705 | -101.9743
7.4741 | -101.9986
7.4778 | -102.0302
7.4815 | -102.0733
7.4851 | -102.1304
7.4888 | -102.1953
7.4924 | -102.2676
7.4961 | -102.3395
7.4997 | -102.4067
7.5034 | -102.4657
7.507 | -102.5171
7.5107 | -102.5496
7.5143 | -102.5671
7.518 | -102.569
7.5217 | -102.552
7.5253 | -102.5115
7.529 | -102.4561
7.5326 | -102.3915
7.5363 | -102.3127
7.5399 | -102.2348
7.5436 | -102.1558
7.5472 | -102.0727
7.5509 | -101.9949
7.5545 | -101.9264
7.5582 | -101.8751
7.5619 | -101.8353
7.5655 | -101.8073
7.5692 | -101.7922
7.5728 | -101.7946
7.5765 | -101.805
7.5801 | -101.8254
7.5838 | -101.8572
7.5874 | -101.895
7.5911 | -101.9402
7.5948 | -101.9883
7.5984 | -102.0338
7.6021 | -102.0803
7.6057 | -102.1208
7.6094 | -102.1615
7.613 | -102.1972
7.6167 | -102.22
7.6203 | -102.2347
7.624 | -102.2496
7.6276 | -102.2566
7.6313 | -102.2539
7.635 | -102.2443
7.6386 | -102.2107
7.6423 | -102.1685
7.6459 | -102.1263
7.6496 | -102.0833
7.6532 | -102.0303
7.6569 | -101.9733
7.6605 | -101.9202
7.6642 | -101.8658
7.6678 | -101.8138
7.6715 | -101.7651
7.6752 | -101.7301
7.6788 | -101.7021
7.6825 | -101.6863
7.6861 | -101.6824
7.6898 | -101.6936
7.6934 | -101.7157
7.6971 | -101.7447
7.7007 | -101.7864
7.7044 | -101.8298
7.7081 | -101.8741
7.7117 | -101.9223
7.7154 | -101.9674
7.719 | -102.0096
7.7227 | -102.0453
7.7263 | -102.0667
7.73 | -102.0724
7.7336 | -102.0647
7.7373 | -102.0402
7.7409 | -102.0021
7.7446 | -101.9577
7.7483 | -101.9092
7.7519 | -101.8572
7.7556 | -101.7886
7.7592 | -101.7135
7.7629 | -101.6404
7.7665 | -101.5712
7.7702 | -101.5034
7.7738 | -101.435
7.7775 | -101.3758
7.7811 | -101.3169
7.7848 | -101.2566
7.7885 | -101.2016
7.7921 | -101.1542
7.7958 | -101.1142
7.7994 | -101.0813
7.8031 | -101.0634
7.8067 | -101.0611
7.8104 | -101.0741
7.814 | -101.0958
7.8177 | -101.1322
7.8214 | -101.1783
7.825 | -101.2345
7.8287 | -101.3026
7.8323 | -101.3792
7.836 | -101.4603
7.8396 | -101.5505
7.8433 | -101.6478
7.8469 | -101.742
7.8506 | -101.8329
7.8542 | -101.9248
7.8579 | -102.0114
7.8616 | -102.0877
7.8652 | -102.1681
7.8689 | -102.2468
7.8725 | -102.3131
7.8762 | -102.3692
7.8798 | -102.4205
7.8835 | -102.4653
7.8871 | -102.4996
7.8908 | -102.5144
7.8944 | -102.5272
7.8981 | -102.5389
7.9018 | -102.5308
7.9054 | -102.5164
7.9091 | -102.4691
7.9127 | -102.4371
7.9164 | -102.4006
7.92 | -102.3618
7.9237 | -102.3151
7.9273 | -102.2672
7.931 | -102.2266
7.9347 | -102.1853
7.9383 | -102.142
7.942 | -102.0859
7.9456 | -102.0357
7.9493 | -101.9893
7.9529 | -101.9478
7.9566 | -101.9099
7.9602 | -101.8701
7.9639 | -101.8331
7.9675 | -101.7971
7.9712 | -101.7607
7.9749 | -101.7246
7.9785 | -101.6859
7.9822 | -101.6489
7.9858 | -101.6192
7.9895 | -101.5947
7.9931 | -101.5745
7.9968 | -101.5659
8.0004 | -101.567
8.0041 | -101.585
8.0077 | -101.6167
8.0114 | -101.6536
8.0151 | -101.7013
8.0187 | -101.7631
8.0224 | -101.8306
8.026 | -101.905
8.0297 | -101.9829
8.0333 | -102.0609
8.037 | -102.1351
8.0406 | -102.203
8.0443 | -102.2744
8.048 | -102.3348
8.0516 | -102.3855
8.0553 | -102.4363
8.0589 | -102.4825
8.0626 | -102.5262
8.0662 | -102.5455
8.0699 | -102.5541
8.0735 | -102.5575
8.0772 | -102.5521
8.0808 | -102.5455
8.0845 | -102.5354
8.0882 | -102.5232
8.0918 | -102.5099
8.0955 | -102.4962
8.0991 | -102.4908
8.1028 | -102.4863
8.1064 | -102.4765
8.1101 | -102.4701
8.1137 | -102.4686
8.1174 | -102.4697
8.121 | -102.4769
8.1247 | -102.4818
8.1284 | -102.4833
8.132 | -102.4876
8.1357 | -102.4942
8.1393 | -102.5043
8.143 | -102.513
8.1466 | -102.5253
8.1503 | -102.5394
8.1539 | -102.5536
8.1576 | -102.5653
8.1613 | -102.5754
8.1649 | -102.5817
8.1686 | -102.5912
8.1722 | -102.5911
8.1759 | -102.5824
8.1795 | -102.5715
8.1832 | -102.5578
8.1868 | -102.5478
8.1905 | -102.5332
8.1941 | -102.5063
8.1978 | -102.4798
8.2015 | -102.4521
8.2051 | -102.4242
8.2088 | -102.3949
8.2124 | -102.3662
8.2161 | -102.3411
8.2197 | -102.3133
8.2234 | -102.2871
8.227 | -102.2664
8.2307 | -102.2454
8.2343 | -102.2245
8.238 | -102.2029
8.2417 | -102.1853
8.2453 | -102.1712
8.249 | -102.1636
8.2526 | -102.159
8.2563 | -102.1569
8.2599 | -102.1586
8.2636 | -102.1659
8.2672 | -102.1784
8.2709 | -102.1836
8.2746 | -102.1886
8.2782 | -102.2028
8.2819 | -102.2156
8.2855 | -102.2244
8.2892 | -102.2288
8.2928 | -102.2354
8.2965 | -102.2368
8.3001 | -102.2342
8.3038 | -102.2284
8.3074 | -102.217
8.3111 | -102.2042
8.3148 | -102.1887
8.3184 | -102.1709
8.3221 | -102.1545
8.3257 | -102.1394
8.3294 | -102.129
8.333 | -102.1188
8.3367 | -102.1093
8.3403 | -102.1086
8.344 | -102.1129
8.3476 | -102.1203
8.3513 | -102.1353
8.355 | -102.1515
8.3586 | -102.1651
8.3623 | -102.1735
8.3659 | -102.1778
8.3696 | -102.1801
8.3732 | -102.1778
8.3769 | -102.1721
8.3805 | -102.1624
8.3842 | -102.1524
8.3879 | -102.1401
8.3915 | -102.1375
8.3952 | -102.1321
8.3988 | -102.1131
8.4025 | -102.0874
8.4061 | -102.0554
8.4098 | -102.0332
8.4134 | -102.005
8.4171 | -101.9753
8.4207 | -101.9398
8.4244 | -101.9008
8.4281 | -101.8654
8.4317 | -101.8307
8.4354 | -101.7989
8.439 | -101.7699
8.4427 | -101.7425
8.4463 | -101.7152
8.45 | -101.6949
8.4536 | -101.6823
8.4573 | -101.6756
8.4609 | -101.6763
8.4646 | -101.6859
8.4683 | -101.7022
8.4719 | -101.7231
8.4756 | -101.7501
8.4792 | -101.7876
8.4829 | -101.8312
8.4865 | -101.8805
8.4902 | -101.9303
8.4938 | -101.9731
8.4975 | -102.011
8.5012 | -102.037
8.5048 | -102.0595
8.5085 | -102.0754
8.5121 | -102.0828
8.5158 | -102.0818
8.5194 | -102.0734
8.5231 | -102.0589
8.5267 | -102.0371
8.5304 | -102.0121
8.534 | -101.9721
8.5377 | -101.9225
8.5414 | -101.8733
8.545 | -101.8285
8.5487 | -101.7796
8.5523 | -101.7233
8.556 | -101.6676
8.5596 | -101.6133
8.5633 | -101.5583
8.5669 | -101.5068
8.5706 | -101.4583
8.5742 | -101.4147
8.5779 | -101.3758
8.5816 | -101.3468
8.5852 | -101.3317
8.5889 | -101.3265
8.5925 | -101.3297
8.5962 | -101.3492
8.5998 | -101.3839
8.6035 | -101.425
8.6071 | -101.4735
8.6108 | -101.5295
8.6145 | -101.5944
8.6181 | -101.6685
8.6218 | -101.7501
8.6254 | -101.8338
8.6291 | -101.9205
8.6327 | -102.0015
8.6364 | -102.0823
8.64 | -102.1607
8.6437 | -102.2326
8.6473 | -102.299
8.651 | -102.3571
8.6547 | -102.4067
8.6583 | -102.4481
8.662 | -102.4792
8.6656 | -102.5062
8.6693 | -102.5192
8.6729 | -102.5208
8.6766 | -102.5142
8.6802 | -102.5039
8.6839 | -102.4921
8.6875 | -102.4791
8.6912 | -102.4662
8.6949 | -102.4506
8.6985 | -102.4356
8.7022 | -102.4261
8.7058 | -102.4231
8.7095 | -102.4244
8.7131 | -102.4217
8.7168 | -102.4249
8.7204 | -102.4322
8.7241 | -102.4396
8.7278 | -102.4466
8.7314 | -102.4555
8.7351 | -102.4616
8.7387 | -102.4636
8.7424 | -102.4564
8.746 | -102.44
8.7497 | -102.4174
8.7533 | -102.3884
8.757 | -102.3537
8.7606 | -102.3147
8.7643 | -102.2656
8.768 | -102.2137
8.7716 | -102.1589
8.7753 | -102.1022
8.7789 | -102.0434
8.7826 | -101.9854
8.7862 | -101.9223
8.7899 | -101.8566
8.7935 | -101.7929
8.7972 | -101.7363
8.8008 | -101.6837
8.8045 | -101.6393
8.8082 | -101.598
8.8118 | -101.5558
8.8155 | -101.5192
8.8191 | -101.4913
8.8228 | -101.4671
8.8264 | -101.4494
8.8301 | -101.4343
8.8337 | -101.4202
8.8374 | -101.4101
8.8411 | -101.4118
8.8447 | -101.4156
8.8484 | -101.4171
8.852 | -101.4241
8.8557 | -101.4324
8.8593 | -101.4375
8.863 | -101.4462
8.8666 | -101.4563
8.8703 | -101.4636
8.8739 | -101.4707
8.8776 | -101.4769
8.8813 | -101.4815
8.8849 | -101.489
8.8886 | -101.4968
8.8922 | -101.5059
8.8959 | -101.5169
8.8995 | -101.5325
8.9032 | -101.553
8.9068 | -101.5795
8.9105 | -101.6076
8.9141 | -101.6475
8.9178 | -101.697
8.9215 | -101.7533
8.9251 | -101.8112
8.9288 | -101.8718
8.9324 | -101.934
8.9361 | -101.9938
8.9397 | -102.057
8.9434 | -102.1235
8.947 | -102.1822
8.9507 | -102.2325
8.9544 | -102.2689
8.958 | -102.2969
8.9617 | -102.2749
8.9653 | -102.2807
8.969 | -102.276
8.9726 | -102.2606
8.9763 | -102.237
8.9799 | -102.2074
8.9836 | -102.1717
8.9872 | -102.1351
8.9909 | -102.0981
8.9946 | -102.0565
8.9982 | -102.0167
9.0019 | -101.9857
9.0055 | -101.9591
9.0092 | -101.9366
9.0128 | -101.9257
9.0165 | -101.9177
9.0201 | -101.9123
9.0238 | -101.916
9.0274 | -101.9268
9.0311 | -101.9435
9.0348 | -101.966
9.0384 | -101.9901
9.0421 | -102.0193
9.0457 | -102.0422
9.0494 | -102.0671
9.053 | -102.0906
9.0567 | -102.111
9.0603 | -102.1363
9.064 | -102.1615
9.0677 | -102.184
9.0713 | -102.2014
9.075 | -102.21
9.0786 | -102.2203
9.0823 | -102.2322
9.0859 | -102.2422
9.0896 | -102.2523
9.0932 | -102.266
9.0969 | -102.2785
9.1005 | -102.2944
9.1042 | -102.3157
9.1079 | -102.3438
9.1115 | -102.3799
9.1152 | -102.4237
9.1188 | -102.4691
9.1225 | -102.5158
9.1261 | -102.5705
9.1298 | -102.6268
9.1334 | -102.6825
9.1371 | -102.735
9.1407 | -102.7868
9.1444 | -102.8346
9.1481 | -102.8742
9.1517 | -102.9115
9.1554 | -102.9466
9.159 | -102.9618
9.1627 | -102.9548
9.1663 | -102.9376
9.17 | -102.9052
9.1736 | -102.8717
9.1773 | -102.8278
9.181 | -102.7811
9.1846 | -102.7192
9.1883 | -102.6536
9.1919 | -102.5881
9.1956 | -102.5143
9.1992 | -102.4369
9.2029 | -102.3637
9.2065 | -102.2924
9.2102 | -102.2243
9.2138 | -102.1617
9.2175 | -102.1078
9.2212 | -102.0656
9.2248 | -102.0323
9.2285 | -102.01
9.2321 | -102.0026
9.2358 | -102.0056
9.2394 | -102.0166
9.2431 | -102.0359
9.2467 | -102.0661
9.2504 | -102.1042
9.254 | -102.1509
9.2577 | -102.2056
9.2614 | -102.2613
9.265 | -102.3192
9.2687 | -102.3763
9.2723 | -102.4346
9.276 | -102.4925
9.2796 | -102.5503
9.2833 | -102.6031
9.2869 | -102.6521
9.2906 | -102.7006
9.2943 | -102.7479
9.2979 | -102.7956
9.3016 | -102.8329
9.3052 | -102.8528
9.3089 | -102.8555
9.3125 | -102.8521
9.3162 | -102.8409
9.3198 | -102.8244
9.3235 | -102.7988
9.3271 | -102.7638
9.3308 | -102.7235
9.3345 | -102.6774
9.3381 | -102.6263
9.3418 | -102.5716
9.3454 | -102.5151
9.3491 | -102.4598
9.3527 | -102.4085
9.3564 | -102.3627
9.36 | -102.323
9.3637 | -102.2926
9.3673 | -102.2708
9.371 | -102.2599
9.3747 | -102.2607
9.3783 | -102.2723
9.382 | -102.2945
9.3856 | -102.3246
9.3893 | -102.3611
9.3929 | -102.4105
9.3966 | -102.4689
9.4002 | -102.5329
9.4039 | -102.6019
9.4076 | -102.6692
9.4112 | -102.7342
9.4149 | -102.7991
9.4185 | -102.8637
9.4222 | -102.9227
9.4258 | -102.9757
9.4295 | -103.0317
9.4331 | -103.0848
9.4368 | -103.1278
9.4404 | -103.1723
9.4441 | -103.2123
9.4478 | -103.245
9.4514 | -103.2724
9.4551 | -103.3007
9.4587 | -103.3253
9.4624 | -103.342
9.466 | -103.3619
9.4697 | -103.3843
9.4733 | -103.4036
9.477 | -103.4172
9.4806 | -103.4276
9.4843 | -103.4304
9.488 | -103.4118
9.4916 | -103.409
9.4953 | -103.4004
9.4989 | -103.3864
9.5026 | -103.368
9.5062 | -103.3492
9.5099 | -103.3334
9.5135 | -103.3152
9.5172 | -103.2938
9.5209 | -103.2751
9.5245 | -103.2606
9.5282 | -103.2472
9.5318 | -103.2314
9.5355 | -103.213
9.5391 | -103.1932
9.5428 | -103.179
9.5464 | -103.1693
9.5501 | -103.1641
9.5537 | -103.1625
9.5574 | -103.1628
9.5611 | -103.1633
9.5647 | -103.1671
9.5684 | -103.1742
9.572 | -103.1818
9.5757 | -103.1935
9.5793 | -103.2101
9.583 | -103.2256
9.5866 | -103.2434
9.5903 | -103.2622
9.5939 | -103.2747
9.5976 | -103.2893
9.6013 | -103.3004
9.6049 | -103.3015
9.6086 | -103.3034
9.6122 | -103.3001
9.6159 | -103.2958
9.6195 | -103.2912
9.6232 | -103.2836
9.6268 | -103.2723
9.6305 | -103.2588
9.6342 | -103.2412
9.6378 | -103.2251
9.6415 | -103.2101
9.6451 | -103.1952
9.6488 | -103.181
9.6524 | -103.1665
9.6561 | -103.1569
9.6597 | -103.1481
9.6634 | -103.1376
9.667 | -103.1336
9.6707 | -103.1359
9.6744 | -103.139
9.678 | -103.1491
9.6817 | -103.1646
9.6853 | -103.1771
9.689 | -103.1932
9.6926 | -103.2136
9.6963 | -103.2333
9.6999 | -103.2555
9.7036 | -103.2819
9.7072 | -103.3068
9.7109 | -103.327
9.7146 | -103.3491
9.7182 | -103.3715
9.7219 | -103.3943
9.7255 | -103.4121
9.7292 | -103.4294
9.7328 | -103.4482
9.7365 | -103.4704
9.7401 | -103.4907
9.7438 | -103.5173
9.7475 | -103.5414
9.7511 | -103.5609
9.7548 | -103.5861
9.7584 | -103.6096
9.7621 | -103.6268
9.7657 | -103.6441
9.7694 | -103.6601
9.773 | -103.6666
9.7767 | -103.6693
9.7803 | -103.6737
9.784 | -103.6754
9.7877 | -103.6749
9.7913 | -103.6692
9.795 | -103.6624
9.7986 | -103.6513
9.8023 | -103.637
9.8059 | -103.6199
9.8096 | -103.5942
9.8132 | -103.56
9.8169 | -103.5186
9.8205 | -103.4856
9.8242 | -103.4492
9.8279 | -103.4112
9.8315 | -103.373
9.8352 | -103.3365
9.8388 | -103.3063
9.8425 | -103.2818
9.8461 | -103.2602
9.8498 | -103.2426
9.8534 | -103.2296
9.8571 | -103.2227
9.8608 | -103.2245
9.8644 | -103.2259
9.8681 | -103.2372
9.8717 | -103.2515
9.8754 | -103.268
9.879 | -103.286
9.8827 | -103.3079
9.8863 | -103.3295
9.89 | -103.3584
9.8936 | -103.3859
9.8973 | -103.412
9.901 | -103.4397
9.9046 | -103.4678
9.9083 | -103.4934
9.9119 | -103.5154
9.9156 | -103.5421
9.9192 | -103.5603
9.9229 | -103.567
9.9265 | -103.5636
9.9302 | -103.5318
9.9338 | -103.4926
9.9375 | -103.4555
9.9412 | -103.4184
9.9448 | -103.3736
9.9485 | -103.3184
9.9521 | -103.2536
9.9558 | -103.18
9.9594 | -103.1037
9.9631 | -103.0111
9.9667 | -102.9187
9.9704 | -102.8288
9.9741 | -102.7395
9.9777 | -102.653
9.9814 | -102.5692
9.985 | -102.4932
9.9887 | -102.4302
9.9923 | -102.3815
9.996 | -102.346
9.9996 | -102.3251
10.0033 | -102.3149
10.0069 | -102.3193
10.0106 | -102.3384
10.0143 | -102.3703
10.0179 | -102.4444
10.0216 | -102.4908
10.0252 | -102.5454
10.0289 | -102.606
10.0325 | -102.6701
10.0362 | -102.7321
10.0398 | -102.7966
10.0435 | -102.8618
10.0471 | -102.9245
10.0508 | -102.9844
10.0545 | -103.0402
10.0581 | -103.0921
10.0618 | -103.1376
10.0654 | -103.1832
10.0691 | -103.2286
10.0727 | -103.2756
10.0764 | -103.3276
10.08 | -103.3693
10.0837 | -103.3943
10.0874 | -103.4113
10.091 | -103.4319
10.0947 | -103.4503
10.0983 | -103.4674
10.102 | -103.4879
10.1056 | -103.5061
10.1093 | -103.5199
10.1129 | -103.5301
10.1166 | -103.5381
10.1202 | -103.5403
10.1239 | -103.5425
10.1276 | -103.5492
10.1312 | -103.5536
10.1349 | -103.5557
10.1385 | -103.5562
10.1422 | -103.563
10.1458 | -103.5687
10.1495 | -103.5777
10.1531 | -103.5927
10.1568 | -103.6074
10.1604 | -103.6225
10.1641 | -103.6388
10.1678 | -103.6561
10.1714 | -103.6764
10.1751 | -103.696
10.1787 | -103.71
10.1824 | -103.7224
10.186 | -103.7348
10.1897 | -103.7459
10.1933 | -103.7559
10.197 | -103.7619
10.2007 | -103.764
10.2043 | -103.7642
10.208 | -103.7616
10.2116 | -103.7528
10.2153 | -103.7396
10.2189 | -103.7223
10.2226 | -103.7026
10.2262 | -103.6873
10.2299 | -103.6631
10.2335 | -103.6394
10.2372 | -103.6176
10.2409 | -103.59
10.2445 | -103.5664
10.2482 | -103.5475
10.2518 | -103.5248
10.2555 | -103.5048
10.2591 | -103.4902
10.2628 | -103.4803
10.2664 | -103.4723
10.2701 | -103.4704
10.2737 | -103.4725
10.2774 | -103.4771
10.2811 | -103.4872
10.2847 | -103.5005
10.2884 | -103.5153
10.292 | -103.5292
10.2957 | -103.5446
10.2993 | -103.5618
10.303 | -103.5789
10.3066 | -103.5986
10.3103 | -103.6165
10.314 | -103.6312
10.3176 | -103.6423
10.3213 | -103.6515
10.3249 | -103.6589
10.3286 | -103.6638
10.3322 | -103.6666
10.3359 | -103.6618
10.3395 | -103.654
10.3432 | -103.6418
10.3468 | -103.6309
10.3505 | -103.6197
10.3542 | -103.6077
10.3578 | -103.594
10.3615 | -103.5826
10.3651 | -103.5732
10.3688 | -103.5628
10.3724 | -103.5534
10.3761 | -103.5451
10.3797 | -103.5439
10.3834 | -103.5395
10.387 | -103.5352
10.3907 | -103.5393
10.3944 | -103.5471
10.398 | -103.5553
10.4017 | -103.566
10.4053 | -103.5801
10.409 | -103.5941
10.4126 | -103.6083
10.4163 | -103.6268
10.4199 | -103.6452
10.4236 | -103.6621
10.4273 | -103.6799
10.4309 | -103.6953
10.4346 | -103.7096
10.4382 | -103.7239
10.4419 | -103.7321
10.4455 | -103.7393
10.4492 | -103.7442
10.4528 | -103.7434
10.4565 | -103.7393
10.4601 | -103.7384
10.4638 | -103.737
10.4675 | -103.7355
10.4711 | -103.7299
10.4748 | -103.7262
10.4784 | -103.724
10.4821 | -103.7215
10.4857 | -103.7234
10.4894 | -103.7245
10.493 | -103.7284
10.4967 | -103.7329
10.5003 | -103.7419
10.504 | -103.7494
10.5077 | -103.7615
10.5113 | -103.7748
10.515 | -103.7892
10.5186 | -103.8083
10.5223 | -103.8299
10.5259 | -103.8504
10.5296 | -103.8729
10.5332 | -103.8963
10.5369 | -103.9202
10.5405 | -103.9408
10.5442 | -103.9453
10.5479 | -103.9569
10.5515 | -103.9683
10.5552 | -103.979
10.5588 | -103.9807
10.5625 | -103.9799
10.5661 | -103.974
10.5698 | -103.9638
10.5734 | -103.95
10.5771 | -103.9347
10.5808 | -103.9144
10.5844 | -103.8953
10.5881 | -103.8771
10.5917 | -103.8561
10.5954 | -103.8373
10.599 | -103.8201
10.6027 | -103.806
10.6063 | -103.7924
10.61 | -103.7788
10.6136 | -103.7669
10.6173 | -103.7576
10.621 | -103.7532
10.6246 | -103.7523
10.6283 | -103.7538
10.6319 | -103.7582
10.6356 | -103.7632
10.6392 | -103.7685
10.6429 | -103.7746
10.6465 | -103.785
10.6502 | -103.7972
10.6539 | -103.8084
10.6575 | -103.8242
10.6612 | -103.8383
10.6648 | -103.8525
10.6685 | -103.8672
10.6721 | -103.8792
10.6758 | -103.8915
10.6794 | -103.9039
10.6831 | -103.9126
10.6867 | -103.9177
10.6904 | -103.9137
10.6941 | -103.9124
10.6977 | -103.907
10.7014 | -103.9009
10.705 | -103.8883
10.7087 | -103.8744
10.7123 | -103.8557
10.716 | -103.8327
10.7196 | -103.8093
10.7233 | -103.7806
10.7269 | -103.7453
10.7306 | -103.7083
10.7343 | -103.6654
10.7379 | -103.621
10.7416 | -103.5732
10.7452 | -103.5241
10.7489 | -103.4764
10.7525 | -103.4325
10.7562 | -103.39
10.7598 | -103.3473
10.7635 | -103.3119
10.7672 | -103.2786
10.7708 | -103.2512
10.7745 | -103.2319
10.7781 | -103.2172
10.7818 | -103.2084
10.7854 | -103.1996
10.7891 | -103.1967
10.7927 | -103.2031
10.7964 | -103.2106
10.8 | -103.2214
10.8037 | -103.2334
10.8074 | -103.2518
10.811 | -103.2738
10.8147 | -103.2965
10.8183 | -103.3235
10.822 | -103.3526
10.8256 | -103.3812
10.8293 | -103.4083
10.8329 | -103.4353
10.8366 | -103.4634
10.8402 | -103.4922
10.8439 | -103.522
10.8476 | -103.5558
10.8512 | -103.5783
10.8549 | -103.5966
10.8585 | -103.608
10.8622 | -103.6201
10.8658 | -103.6264
10.8695 | -103.6295
10.8731 | -103.6312
10.8768 | -103.626
10.8804 | -103.6224
10.8841 | -103.6151
10.8878 | -103.6056
10.8914 | -103.5944
10.8951 | -103.581
10.8987 | -103.5667
10.9024 | -103.5513
10.906 | -103.5383
10.9097 | -103.526
10.9133 | -103.5153
10.917 | -103.509
10.9207 | -103.4992
10.9243 | -103.4867
10.928 | -103.4855
10.9316 | -103.4882
10.9353 | -103.496
10.9389 | -103.5075
10.9426 | -103.5254
10.9462 | -103.5475
10.9499 | -103.5708
10.9535 | -103.5984
10.9572 | -103.6275
10.9609 | -103.6563
10.9645 | -103.6811
10.9682 | -103.7067
10.9718 | -103.7301
10.9755 | -103.7518
10.9791 | -103.7726
10.9828 | -103.7882
10.9864 | -103.8021
10.9901 | -103.8119
10.9938 | -103.8226
10.9974 | -103.8314
11.0011 | -103.837
11.0047 | -103.8387
11.0084 | -103.8403
11.012 | -103.8443
11.0157 | -103.8475
11.0193 | -103.8449
11.023 | -103.8437
11.0266 | -103.8417
11.0303 | -103.835
11.034 | -103.8265
11.0376 | -103.8119
11.0413 | -103.7959
11.0449 | -103.7764
11.0486 | -103.7576
11.0522 | -103.7389
11.0559 | -103.7182
11.0595 | -103.6962
11.0632 | -103.6684
11.0668 | -103.6433
11.0705 | -103.632
11.0742 | -103.6192
11.0778 | -103.6086
11.0815 | -103.6026
11.0851 | -103.6044
11.0888 | -103.5978
11.0924 | -103.5946
11.0961 | -103.5919
11.0997 | -103.5965
11.1034 | -103.6068
11.107 | -103.6181
11.1107 | -103.6298
11.1144 | -103.6467
11.118 | -103.6659
11.1217 | -103.6862
11.1253 | -103.7084
11.129 | -103.7294
11.1326 | -103.7534
11.1363 | -103.7759
11.1399 | -103.7964
11.1436 | -103.819
11.1473 | -103.8431
11.1509 | -103.8658
11.1546 | -103.8839
11.1582 | -103.8984
11.1619 | -103.9139
11.1655 | -103.9288
11.1692 | -103.9446
11.1728 | -103.9557
11.1765 | -103.9633
11.1801 | -103.9642
11.1838 | -103.9636
11.1875 | -103.9683
11.1911 | -103.9777
11.1948 | -103.9839
11.1984 | -103.9896
11.2021 | -103.9931
11.2057 | -103.9977
11.2094 | -104.0011
11.213 | -104.0026
11.2167 | -104.0036
11.2203 | -104.0052
11.224 | -104.0092
11.2277 | -104.0067
11.2313 | -104.0013
11.235 | -103.9944
11.2386 | -103.9858
11.2423 | -103.9775
11.2459 | -103.9702
11.2496 | -103.9609
11.2532 | -103.9472
11.2569 | -103.9311
11.2606 | -103.9162
11.2642 | -103.8978
11.2679 | -103.8825
11.2715 | -103.8611
11.2752 | -103.838
11.2788 | -103.8137
11.2825 | -103.7874
11.2861 | -103.7593
11.2898 | -103.7286
11.2934 | -103.6971
11.2971 | -103.6625
11.3008 | -103.6265
11.3044 | -103.5911
11.3081 | -103.5568
11.3117 | -103.5195
11.3154 | -103.4779
11.319 | -103.4344
11.3227 | -103.3944
11.3263 | -103.3552
11.33 | -103.3153
11.3337 | -103.2713
11.3373 | -103.2289
11.341 | -103.1929
11.3446 | -103.1621
11.3483 | -103.1322
11.3519 | -103.1038
11.3556 | -103.0767
11.3592 | -103.052
11.3629 | -103.0317
11.3665 | -103.021
11.3702 | -103.0136
11.3739 | -103.0093
11.3775 | -103.0095
11.3812 | -103.0128
11.3848 | -103.0199
11.3885 | -103.0298
11.3921 | -103.0429
11.3958 | -103.0563
11.3994 | -103.0719
11.4031 | -103.0878
11.4067 | -103.1097
11.4104 | -103.1349
11.4141 | -103.1585
11.4177 | -103.1785
11.4214 | -103.1995
11.425 | -103.22
11.4287 | -103.2383
11.4323 | -103.2503
11.436 | -103.2649
11.4396 | -103.2788
11.4433 | -103.2902
11.4469 | -103.2994
11.4506 | -103.302
11.4543 | -103.299
11.4579 | -103.2856
11.4616 | -103.2676
11.4652 | -103.2496
11.4689 | -103.2275
11.4725 | -103.2004
11.4762 | -103.1722
11.4798 | -103.144
11.4835 | -103.1178
11.4872 | -103.0879
11.4908 | -103.0554
11.4945 | -103.0261
11.4981 | -103.0022
11.5018 | -102.9799
11.5054 | -102.9569
11.5091 | -102.939
11.5127 | -102.9247
11.5164 | -102.9145
11.52 | -102.9104
11.5237 | -102.9163
11.5274 | -102.9241
11.531 | -102.9317
11.5347 | -102.9432
11.5383 | -102.9648
11.542 | -102.9921
11.5456 | -103.0264
11.5493 | -103.0615
11.5529 | -103.1003
11.5566 | -103.1436
11.5602 | -103.189
11.5639 | -103.2366
11.5676 | -103.2857
11.5712 | -103.3281
11.5749 | -103.3698
11.5785 | -103.4134
11.5822 | -103.456
11.5858 | -103.4959
11.5895 | -103.5341
11.5931 | -103.5675
11.5968 | -103.5842
11.6005 | -103.6134
11.6041 | -103.6374
11.6078 | -103.6599
11.6114 | -103.6746
11.6151 | -103.6852
11.6187 | -103.7019
11.6224 | -103.7185
11.626 | -103.7362
11.6297 | -103.7475
11.6333 | -103.7493
11.637 | -103.7483
11.6407 | -103.7477
11.6443 | -103.7436
11.648 | -103.7349
11.6516 | -103.7253
11.6553 | -103.7166
11.6589 | -103.7072
11.6626 | -103.6949
11.6662 | -103.6803
11.6699 | -103.6666
11.6736 | -103.6532
11.6772 | -103.6354
11.6809 | -103.6173
11.6845 | -103.6012
11.6882 | -103.5859
11.6918 | -103.5684
11.6955 | -103.5568
11.6991 | -103.5484
11.7028 | -103.5428
11.7064 | -103.5352
11.7101 | -103.5288
11.7138 | -103.5265
11.7174 | -103.5287
11.7211 | -103.5364
11.7247 | -103.5458
11.7284 | -103.5611
11.732 | -103.5775
11.7357 | -103.5952
11.7393 | -103.6197
11.743 | -103.6452
11.7466 | -103.6703
11.7503 | -103.6982
11.754 | -103.7213
11.7576 | -103.743
11.7613 | -103.7651
11.7649 | -103.7851
11.7686 | -103.8029
11.7722 | -103.8188
11.7759 | -103.8324
11.7795 | -103.8402
11.7832 | -103.8496
11.7868 | -103.8581
11.7905 | -103.861
11.7942 | -103.8622
11.7978 | -103.8641
11.8015 | -103.8658
11.8051 | -103.8649
11.8088 | -103.8601
11.8124 | -103.8543
11.8161 | -103.8479
11.8197 | -103.8447
11.8234 | -103.8363
11.8271 | -103.8252
11.8307 | -103.8146
11.8344 | -103.8077
11.838 | -103.7989
11.8417 | -103.7875
11.8453 | -103.7766
11.849 | -103.7708
11.8526 | -103.7658
11.8563 | -103.7608
11.8599 | -103.7567
11.8636 | -103.754
11.8673 | -103.7508
11.8709 | -103.7478
11.8746 | -103.7464
11.8782 | -103.7436
11.8819 | -103.7448
11.8855 | -103.7428
11.8892 | -103.7435
11.8928 | -103.7431
11.8965 | -103.7417
11.9002 | -103.741
11.9038 | -103.7418
11.9075 | -103.7454
11.9111 | -103.7431
11.9148 | -103.7238
11.9184 | -103.7113
11.9221 | -103.7065
11.9257 | -103.7009
11.9294 | -103.6943
11.933 | -103.6882
11.9367 | -103.6822
11.9404 | -103.6739
11.944 | -103.6642
11.9477 | -103.6539
11.9513 | -103.6425
11.955 | -103.6298
11.9586 | -103.612
11.9623 | -103.5921
11.9659 | -103.5741
11.9696 | -103.5546
11.9732 | -103.5345
11.9769 | -103.5157
11.9806 | -103.4991
11.9842 | -103.4838
11.9879 | -103.4713
11.9915 | -103.4593
11.9952 | -103.452
11.9988 | -103.4491
12.0025 | -103.448
12.0061 | -103.4464
12.0098 | -103.4541
12.0135 | -103.4666
12.0171 | -103.4803
12.0208 | -103.4972
12.0244 | -103.5205
12.0281 | -103.5474
12.0317 | -103.5744
12.0354 | -103.6056
12.039 | -103.6419
12.0427 | -103.6802
12.0463 | -103.7177
12.05 | -103.7577
12.0537 | -103.7966
12.0573 | -103.8381
12.061 | -103.8805
12.0646 | -103.9206
12.0683 | -103.9603
12.0719 | -103.995
12.0756 | -104.0284
12.0792 | -104.0608
12.0829 | -104.0947
12.0865 | -104.1367
12.0902 | -104.1652
12.0939 | -104.1775
12.0975 | -104.1875
12.1012 | -104.1964
12.1048 | -104.2
12.1085 | -104.2041
12.1121 | -104.2037
12.1158 | -104.2023
12.1194 | -104.2008
12.1231 | -104.1993
12.1267 | -104.1952
12.1304 | -104.1916
12.1341 | -104.1891
12.1377 | -104.1855
12.1414 | -104.1782
12.145 | -104.1728
12.1487 | -104.1727
12.1523 | -104.1712
12.156 | -104.1699
12.1596 | -104.1662
12.1633 | -104.1623
12.167 | -104.1593
12.1706 | -104.1573
12.1743 | -104.1555
12.1779 | -104.1543
12.1816 | -104.1507
12.1852 | -104.1427
12.1889 | -104.1415
12.1925 | -104.1393
12.1962 | -104.1352
12.1998 | -104.1321
12.2035 | -104.1303
12.2072 | -104.1328
12.2108 | -104.1336
12.2145 | -104.1277
12.2181 | -104.1191
12.2218 | -104.1106
12.2254 | -104.1054
12.2291 | -104.1008
12.2327 | -104.0961
12.2364 | -104.0878
12.2401 | -104.08
12.2437 | -104.0742
12.2474 | -104.0688
12.251 | -104.0621
12.2547 | -104.0502
12.2583 | -104.0367
12.262 | -104.0241
12.2656 | -104.0045
12.2693 | -103.9863
12.2729 | -103.9689
12.2766 | -103.9488
12.2803 | -103.9299
12.2839 | -103.9093
12.2876 | -103.887
12.2912 | -103.8659
12.2949 | -103.8462
12.2985 | -103.8293
12.3022 | -103.8132
12.3058 | -103.7965
12.3095 | -103.7827
12.3131 | -103.7724
12.3168 | -103.7666
12.3205 | -103.7632
12.3241 | -103.7583
12.3278 | -103.7539
12.3314 | -103.7536
12.3351 | -103.7522
12.3387 | -103.7533
12.3424 | -103.7584
12.346 | -103.7645
12.3497 | -103.7702
12.3534 | -103.7778
12.357 | -103.7863
12.3607 | -103.7977
12.3643 | -103.8107
12.368 | -103.8176
12.3716 | -103.8267
12.3753 | -103.8368
12.3789 | -103.8428
12.3826 | -103.8481
12.3862 | -103.8543
12.3899 | -103.8599
12.3936 | -103.8694
12.3972 | -103.8826
12.4009 | -103.891
12.4045 | -103.8908
12.4082 | -103.8897
12.4118 | -103.8836
12.4155 | -103.8792
12.4191 | -103.8738
12.4228 | -103.8671
12.4264 | -103.8575
12.4301 | -103.8433
12.4338 | -103.8268
12.4374 | -103.8068
12.4411 | -103.7824
12.4447 | -103.7592
12.4484 | -103.7322
12.452 | -103.6989
12.4557 | -103.6654
12.4593 | -103.6295
12.463 | -103.5924
12.4666 | -103.558
12.4703 | -103.5237
12.474 | -103.4899
12.4776 | -103.4562
12.4813 | -103.425
12.4849 | -103.3992
12.4886 | -103.3771
12.4922 | -103.3595
12.4959 | -103.3415
12.4995 | -103.3286
12.5032 | -103.3242
12.5069 | -103.3263
12.5105 | -103.3342
12.5142 | -103.3455
12.5178 | -103.3597
12.5215 | -103.3778
12.5251 | -103.3999
12.5288 | -103.4248
12.5324 | -103.4488
12.5361 | -103.4743
12.5397 | -103.5022
12.5434 | -103.5281
12.5471 | -103.5563
12.5507 | -103.5871
12.5544 | -103.6162
12.558 | -103.6439
12.5617 | -103.6697
12.5653 | -103.6962
12.569 | -103.7228
12.5726 | -103.7463
12.5763 | -103.7634
12.58 | -103.7822
12.5836 | -103.8033
12.5873 | -103.8195
12.5909 | -103.8351
12.5946 | -103.8459
12.5982 | -103.8584
12.6019 | -103.8674
12.6055 | -103.8749
12.6092 | -103.883
12.6128 | -103.8922
12.6165 | -103.9
12.6202 | -103.9069
12.6238 | -103.9089
12.6275 | -103.9114
12.6311 | -103.9128
12.6348 | -103.9128
12.6384 | -103.9142
12.6421 | -103.9144
12.6457 | -103.9167
12.6494 | -103.9171
12.653 | -103.9099
12.6567 | -103.9045
12.6604 | -103.8971
12.664 | -103.8927
12.6677 | -103.8883
12.6713 | -103.8809
12.675 | -103.8684
12.6786 | -103.8595
12.6823 | -103.8511
12.6859 | -103.8415
12.6896 | -103.8277
12.6933 | -103.8117
12.6969 | -103.798
12.7006 | -103.7861
12.7042 | -103.7742
12.7079 | -103.7607
12.7115 | -103.7476
12.7152 | -103.7375
12.7188 | -103.7276
12.7225 | -103.7163
12.7261 | -103.7096
12.7298 | -103.702
12.7335 | -103.6926
12.7371 | -103.6876
12.7408 | -103.6848
12.7444 | -103.6842
12.7481 | -103.6817
12.7517 | -103.6809
12.7554 | -103.682
12.759 | -103.6881
12.7627 | -103.6949
12.7663 | -103.7005
12.77 | -103.7072
12.7737 | -103.715
12.7773 | -103.7225
12.781 | -103.7242
12.7846 | -103.7286
12.7883 | -103.7375
12.7919 | -103.7447
12.7956 | -103.7492
12.7992 | -103.7535
12.8029 | -103.7586
12.8065 | -103.7635
12.8102 | -103.7629
12.8139 | -103.7629
12.8175 | -103.7607
12.8212 | -103.7581
12.8248 | -103.7569
12.8285 | -103.7499
12.8321 | -103.743
12.8358 | -103.7385
12.8394 | -103.7326
12.8431 | -103.7294
12.8468 | -103.727
12.8504 | -103.7219
12.8541 | -103.7139
12.8577 | -103.7081
12.8614 | -103.7048
12.865 | -103.6962
12.8687 | -103.687
12.8723 | -103.6776
12.876 | -103.6701
12.8796 | -103.664
12.8833 | -103.6606
12.887 | -103.6546
12.8906 | -103.6461
12.8943 | -103.6425
12.8979 | -103.6387
12.9016 | -103.6342
12.9052 | -103.625
12.9089 | -103.6078
12.9125 | -103.5967
12.9162 | -103.5868
12.9198 | -103.576
12.9235 | -103.5632
12.9272 | -103.5505
12.9308 | -103.5371
12.9345 | -103.5262
12.9381 | -103.5117
12.9418 | -103.4954
12.9454 | -103.4801
12.9491 | -103.4649
12.9527 | -103.45
12.9564 | -103.4353
12.9601 | -103.418
12.9637 | -103.3996
12.9674 | -103.3827
12.971 | -103.3622
12.9747 | -103.341
12.9783 | -103.3245
12.982 | -103.3095
12.9856 | -103.2918
12.9893 | -103.2777
12.9929 | -103.2656
12.9966 | -103.2549
13.0003 | -103.2461
13.0039 | -103.242
13.0076 | -103.2389
13.0112 | -103.239
13.0149 | -103.242
13.0185 | -103.2468
13.0222 | -103.2517
13.0258 | -103.2559
13.0295 | -103.2579
13.0331 | -103.2602
13.0368 | -103.2626
13.0405 | -103.2638
13.0441 | -103.2619
13.0478 | -103.2614
13.0514 | -103.262
13.0551 | -103.2612
13.0587 | -103.2524
13.0624 | -103.2413
13.066 | -103.2304
13.0697 | -103.2187
13.0734 | -103.2087
13.077 | -103.2
13.0807 | -103.1884
13.0843 | -103.1764
13.088 | -103.1673
13.0916 | -103.1598
13.0953 | -103.158
13.0989 | -103.1459
13.1026 | -103.1368
13.1062 | -103.1277
13.1099 | -103.1188
13.1136 | -103.113
13.1172 | -103.1058
13.1209 | -103.0972
13.1245 | -103.0886
13.1282 | -103.081
13.1318 | -103.0735
13.1355 | -103.0625
13.1391 | -103.0539
13.1428 | -103.0459
13.1464 | -103.0381
13.1501 | -103.0335
13.1538 | -103.0274
13.1574 | -103.0242
13.1611 | -103.0201
13.1647 | -103.0147
13.1684 | -103.0097
13.172 | -103.0067
13.1757 | -103.0005
13.1793 | -102.9914
13.183 | -102.9844
13.1867 | -102.9783
13.1903 | -102.9696
13.194 | -102.9565
13.1976 | -102.9413
13.2013 | -102.9248
13.2049 | -102.9104
13.2086 | -102.8939
13.2122 | -102.878
13.2159 | -102.8627
13.2195 | -102.8463
13.2232 | -102.8297
13.2269 | -102.8119
13.2305 | -102.7931
13.2342 | -102.7735
13.2378 | -102.7534
13.2415 | -102.7345
13.2451 | -102.7126
13.2488 | -102.6904
13.2524 | -102.6653
13.2561 | -102.6403
13.2597 | -102.6113
13.2634 | -102.5753
13.2671 | -102.5343
13.2707 | -102.4974
13.2744 | -102.4603
13.278 | -102.4209
13.2817 | -102.3778
13.2853 | -102.3316
13.289 | -102.2854
13.2926 | -102.235
13.2963 | -102.1854
13.3 | -102.136
13.3036 | -102.0848
13.3073 | -102.0347
13.3109 | -101.9857
13.3146 | -101.9401
13.3182 | -101.894
13.3219 | -101.8513
13.3255 | -101.8151
13.3292 | -101.7802
13.3328 | -101.751
13.3365 | -101.7272
13.3402 | -101.7072
13.3438 | -101.6951
13.3475 | -101.6891
13.3511 | -101.6899
13.3548 | -101.6989
13.3584 | -101.716
13.3621 | -101.7425
13.3657 | -101.7781
13.3694 | -101.8218
13.373 | -101.872
13.3767 | -101.9307
13.3804 | -101.9994
13.384 | -102.0771
13.3877 | -102.1625
13.3913 | -102.2538
13.395 | -102.3506
13.3986 | -102.4519
13.4023 | -102.558
13.4059 | -102.6682
13.4096 | -102.7828
13.4133 | -102.8961
13.4169 | -103.0096
13.4206 | -103.1229
13.4242 | -103.2364
13.4279 | -103.3469
13.4315 | -103.4516
13.4352 | -103.5503
13.4388 | -103.6458
13.4425 | -103.7362
13.4461 | -103.8202
13.4498 | -103.8898
13.4535 | -103.9512
13.4571 | -104.0103
13.4608 | -104.0618
13.4644 | -104.0947
13.4681 | -104.118
13.4717 | -104.1316
13.4754 | -104.1457
13.479 | -104.1502
13.4827 | -104.1465
13.4863 | -104.1389
13.49 | -104.1255
13.4937 | -104.1071
13.4973 | -104.0866
13.501 | -104.0646
13.5046 | -104.0364
13.5083 | -104.0079
13.5119 | -103.9831
13.5156 | -103.9559
13.5192 | -103.927
13.5229 | -103.9016
13.5266 | -103.8719
13.5302 | -103.8482
13.5339 | -103.8258
13.5375 | -103.8016
13.5412 | -103.777
13.5448 | -103.7534
13.5485 | -103.7286
13.5521 | -103.7046
13.5558 | -103.6801
13.5594 | -103.6571
13.5631 | -103.6382
13.5668 | -103.6198
13.5704 | -103.5993
13.5741 | -103.5779
13.5777 | -103.5568
13.5814 | -103.539
13.585 | -103.5259
13.5887 | -103.5151
13.5923 | -103.5058
13.596 | -103.4978
13.5996 | -103.4913
13.6033 | -103.4873
13.607 | -103.4889
13.6106 | -103.4905
13.6143 | -103.4941
13.6179 | -103.4999
13.6216 | -103.5085
13.6252 | -103.5169
13.6289 | -103.5262
13.6325 | -103.5382
13.6362 | -103.5526
13.6399 | -103.5663
13.6435 | -103.5811
13.6472 | -103.5954
13.6508 | -103.6073
13.6545 | -103.6205
13.6581 | -103.6339
13.6618 | -103.6473
13.6654 | -103.6598
13.6691 | -103.6707
13.6727 | -103.6843
13.6764 | -103.6976
13.6801 | -103.7108
13.6837 | -103.721
13.6874 | -103.7343
13.691 | -103.7461
13.6947 | -103.7567
13.6983 | -103.7699
13.702 | -103.7781
13.7056 | -103.7803
13.7093 | -103.7872
13.7129 | -103.7942
13.7166 | -103.8021
13.7203 | -103.8076
13.7239 | -103.8157
13.7276 | -103.8209
13.7312 | -103.8269
13.7349 | -103.8327
13.7385 | -103.8369
13.7422 | -103.8389
13.7458 | -103.8431
13.7495 | -103.8465
13.7532 | -103.8494
13.7568 | -103.8492
13.7605 | -103.8482
13.7641 | -103.8504
13.7678 | -103.8559
13.7714 | -103.858
13.7751 | -103.8581
13.7787 | -103.8571
13.7824 | -103.855
13.786 | -103.8511
13.7897 | -103.8488
13.7934 | -103.8455
13.797 | -103.8402
13.8007 | -103.8328
13.8043 | -103.8266
13.808 | -103.8203
13.8116 | -103.8136
13.8153 | -103.8093
13.8189 | -103.8047
13.8226 | -103.8
13.8262 | -103.7973
13.8299 | -103.7966
13.8336 | -103.7971
13.8372 | -103.7964
13.8409 | -103.7955
13.8445 | -103.7976
13.8482 | -103.8011
13.8518 | -103.8056
13.8555 | -103.8084
13.8591 | -103.8114
13.8628 | -103.8152
13.8665 | -103.8191
13.8701 | -103.8214
13.8738 | -103.8254
13.8774 | -103.8286
13.8811 | -103.8313
13.8847 | -103.8342
13.8884 | -103.8376
13.892 | -103.8409
13.8957 | -103.8449
13.8993 | -103.8458
13.903 | -103.8433
13.9067 | -103.8408
13.9103 | -103.8375
13.914 | -103.83
13.9176 | -103.824
13.9213 | -103.8199
13.9249 | -103.8153
13.9286 | -103.8089
13.9322 | -103.8018
13.9359 | -103.7975
13.9395 | -103.7929
13.9432 | -103.7899
13.9469 | -103.7847
13.9505 | -103.7799
13.9542 | -103.7748
13.9578 | -103.7707
13.9615 | -103.7664
13.9651 | -103.7635
13.9688 | -103.7629
13.9724 | -103.7642
13.9761 | -103.769
13.9798 | -103.7743
13.9834 | -103.7826
13.9871 | -103.7937
13.9907 | -103.8038
13.9944 | -103.8163
13.998 | -103.8321
14.0017 | -103.8483
14.0053 | -103.8657
14.009 | -103.8872
14.0126 | -103.9083
14.0163 | -103.9271
14.02 | -103.9472
14.0236 | -103.9694
14.0273 | -103.9916
14.0309 | -104.0107
14.0346 | -104.032
14.0382 | -104.0581
14.0419 | -104.086
14.0455 | -104.1118
14.0492 | -104.1374
14.0528 | -104.1631
14.0565 | -104.1864
14.0602 | -104.2077
14.0638 | -104.2276
14.0675 | -104.2456
14.0711 | -104.2668
14.0748 | -104.2867
14.0784 | -104.3075
14.0821 | -104.3282
14.0857 | -104.3432
14.0894 | -104.3564
14.0931 | -104.3684
14.0967 | -104.3766
14.1004 | -104.3891
14.104 | -104.4016
14.1077 | -104.4064
14.1113 | -104.4102
14.115 | -104.4122
14.1186 | -104.4105
14.1223 | -104.409
14.1259 | -104.4062
14.1296 | -104.4034
14.1333 | -104.3988
14.1369 | -104.3921
14.1406 | -104.3864
14.1442 | -104.3804
14.1479 | -104.3764
14.1515 | -104.3701
14.1552 | -104.3668
14.1588 | -104.3635
14.1625 | -104.3581
14.1661 | -104.3552
14.1698 | -104.3524
14.1735 | -104.3528
14.1771 | -104.3525
14.1808 | -104.355
14.1844 | -104.3572
14.1881 | -104.3616
14.1917 | -104.3703
14.1954 | -104.3794
14.199 | -104.3884
14.2027 | -104.4008
14.2064 | -104.4134
14.21 | -104.4264
14.2137 | -104.4422
14.2173 | -104.4581
14.221 | -104.4769
14.2246 | -104.4974
14.2283 | -104.5185
14.2319 | -104.5395
14.2356 | -104.5628
14.2392 | -104.5841
14.2429 | -104.6045
14.2466 | -104.6247
14.2502 | -104.6471
14.2539 | -104.6699
14.2575 | -104.6912
14.2612 | -104.7156
14.2648 | -104.7372
14.2685 | -104.759
14.2721 | -104.781
14.2758 | -104.7981
14.2794 | -104.8115
14.2831 | -104.8233
14.2868 | -104.8342
14.2904 | -104.8452
14.2941 | -104.8524
14.2977 | -104.8572
14.3014 | -104.8626
14.305 | -104.8628
14.3087 | -104.8613
14.3123 | -104.8586
14.316 | -104.8556
14.3197 | -104.8504
14.3233 | -104.8414
14.327 | -104.8311
14.3306 | -104.8224
14.3343 | -104.813
14.3379 | -104.8038
14.3416 | -104.7973
14.3452 | -104.7919
14.3489 | -104.7865
14.3525 | -104.779
14.3562 | -104.7751
14.3599 | -104.7695
14.3635 | -104.7629
14.3672 | -104.7579
14.3708 | -104.75
14.3745 | -104.7446
14.3781 | -104.7412
14.3818 | -104.7383
14.3854 | -104.7364
14.3891 | -104.7358
14.3927 | -104.7328
14.3964 | -104.7276
14.4001 | -104.724
14.4037 | -104.7179
14.4074 | -104.7117
14.411 | -104.7047
14.4147 | -104.696
14.4183 | -104.6861
14.422 | -104.6791
14.4256 | -104.6698
14.4293 | -104.6595
14.433 | -104.6483
14.4366 | -104.6411
14.4403 | -104.6352
14.4439 | -104.6258
14.4476 | -104.6168
14.4512 | -104.6095
14.4549 | -104.6033
14.4585 | -104.5965
14.4622 | -104.5919
14.4658 | -104.589
14.4695 | -104.5865
14.4732 | -104.5855
14.4768 | -104.5829
14.4805 | -104.5826
14.4841 | -104.5846
14.4878 | -104.5875
14.4914 | -104.5888
14.4951 | -104.5932
14.4987 | -104.597
14.5024 | -104.6013
14.506 | -104.6064
14.5097 | -104.6116
14.5134 | -104.6191
14.517 | -104.6251
14.5207 | -104.6298
14.5243 | -104.633
14.528 | -104.64
14.5316 | -104.6499
14.5353 | -104.6574
14.5389 | -104.664
14.5426 | -104.6703
14.5463 | -104.6762
14.5499 | -104.6782
14.5536 | -104.6758
14.5572 | -104.6719
14.5609 | -104.6701
14.5645 | -104.6698
14.5682 | -104.668
14.5718 | -104.6671
14.5755 | -104.6663
14.5791 | -104.6636
14.5828 | -104.657
14.5865 | -104.6498
14.5901 | -104.6444
14.5938 | -104.6371
14.5974 | -104.6326
14.6011 | -104.6283
14.6047 | -104.625
14.6084 | -104.6231
14.612 | -104.6237
14.6157 | -104.6244
14.6193 | -104.6254
14.623 | -104.6266
14.6267 | -104.6287
14.6303 | -104.6291
14.634 | -104.6309
14.6376 | -104.6339
14.6413 | -104.6363
14.6449 | -104.64
14.6486 | -104.6441
14.6522 | -104.6515
14.6559 | -104.6572
14.6596 | -104.6612
14.6632 | -104.665
14.6669 | -104.6708
14.6705 | -104.6756
14.6742 | -104.6808
14.6778 | -104.6871
14.6815 | -104.6939
14.6851 | -104.6989
14.6888 | -104.7007
14.6924 | -104.7037
14.6961 | -104.7078
14.6998 | -104.7163
14.7034 | -104.7225
14.7071 | -104.7266
14.7107 | -104.7299
14.7144 | -104.7296
14.718 | -104.7305
14.7217 | -104.7298
14.7253 | -104.7305
14.729 | -104.7308
14.7326 | -104.7302
14.7363 | -104.7264
14.74 | -104.7227
14.7436 | -104.7164
14.7473 | -104.7118
14.7509 | -104.7078
14.7546 | -104.7037
14.7582 | -104.7021
14.7619 | -104.7022
14.7655 | -104.7059
14.7692 | -104.7105
14.7729 | -104.7114
14.7765 | -104.7077
14.7802 | -104.7059
14.7838 | -104.7049
14.7875 | -104.7053
14.7911 | -104.7068
14.7948 | -104.7084
14.7984 | -104.7099
14.8021 | -104.712
14.8057 | -104.7163
14.8094 | -104.7187
14.8131 | -104.7245
14.8167 | -104.7324
14.8204 | -104.7379
14.824 | -104.7415
14.8277 | -104.7451
14.8313 | -104.7505
14.835 | -104.7576
14.8386 | -104.7646
14.8423 | -104.7732
14.8459 | -104.7802
14.8496 | -104.7847
14.8533 | -104.7918
14.8569 | -104.7969
14.8606 | -104.8013
14.8642 | -104.8054
14.8679 | -104.8118
14.8715 | -104.8164
14.8752 | -104.8211
14.8788 | -104.8245
14.8825 | -104.8253
14.8862 | -104.8266
14.8898 | -104.8303
14.8935 | -104.8277
14.8971 | -104.8259
14.9008 | -104.8241
14.9044 | -104.8208
14.9081 | -104.8171
14.9117 | -104.8091
14.9154 | -104.803
14.919 | -104.8002
14.9227 | -104.7974
14.9264 | -104.794
14.93 | -104.7831
14.9337 | -104.7713
14.9373 | -104.7585
14.941 | -104.7472
14.9446 | -104.737
14.9483 | -104.7254
14.9519 | -104.7093
14.9556 | -104.6957
14.9592 | -104.6806
14.9629 | -104.6655
14.9666 | -104.6472
14.9702 | -104.6303
14.9739 | -104.6145
14.9775 | -104.5976
14.9812 | -104.5812
14.9848 | -104.5656
14.9885 | -104.5527
14.9921 | -104.5439
14.9958 | -104.535
14.9995 | -104.5281
15.0031 | -104.5241
15.0068 | -104.5203
15.0104 | -104.5166
15.0141 | -104.5144
15.0177 | -104.5149
15.0214 | -104.5147
15.025 | -104.5151
15.0287 | -104.5168
15.0323 | -104.5209
15.036 | -104.5278
15.0397 | -104.5376
15.0433 | -104.5493
15.047 | -104.5621
15.0506 | -104.5723
15.0543 | -104.5832
15.0579 | -104.5932
15.0616 | -104.6027
15.0652 | -104.6156
15.0689 | -104.6261
15.0725 | -104.6371
15.0762 | -104.6472
15.0799 | -104.6567
15.0835 | -104.6653
15.0872 | -104.6739
15.0908 | -104.6813
15.0945 | -104.6874
15.0981 | -104.6939
15.1018 | -104.6977
15.1054 | -104.6971
15.1091 | -104.6985
15.1128 | -104.6986
15.1164 | -104.6942
15.1201 | -104.6862
15.1237 | -104.6793
15.1274 | -104.6745
15.131 | -104.6676
15.1347 | -104.6558
15.1383 | -104.6457
15.142 | -104.6374
15.1456 | -104.6328
15.1493 | -104.6251
15.153 | -104.6133
15.1566 | -104.6015
15.1603 | -104.5895
15.1639 | -104.5812
15.1676 | -104.5765
15.1712 | -104.5685
15.1749 | -104.5611
15.1785 | -104.555
15.1822 | -104.5541
15.1858 | -104.5502
15.1895 | -104.5451
15.1932 | -104.5433
15.1968 | -104.5425
15.2005 | -104.54
15.2041 | -104.5403
15.2078 | -104.5412
15.2114 | -104.5377
15.2151 | -104.5363
15.2187 | -104.5369
15.2224 | -104.5384
15.2261 | -104.5454
15.2297 | -104.552
15.2334 | -104.555
15.237 | -104.5596
15.2407 | -104.5691
15.2443 | -104.5761
15.248 | -104.5828
15.2516 | -104.5908
15.2553 | -104.596
15.2589 | -104.5992
15.2626 | -104.6031
15.2663 | -104.611
15.2699 | -104.6167
15.2736 | -104.6216
15.2772 | -104.6247
15.2809 | -104.6246
15.2845 | -104.6252
15.2882 | -104.6211
15.2918 | -104.6193
15.2955 | -104.6163
15.2991 | -104.6108
15.3028 | -104.6035
15.3065 | -104.5968
15.3101 | -104.5905
15.3138 | -104.5826
15.3174 | -104.575
15.3211 | -104.5673
15.3247 | -104.5586
15.3284 | -104.5491
15.332 | -104.538
15.3357 | -104.5269
15.3394 | -104.5162
15.343 | -104.5037
15.3467 | -104.4901
15.3503 | -104.4793
15.354 | -104.4704
15.3576 | -104.4614
15.3613 | -104.454
15.3649 | -104.4462
15.3686 | -104.4409
15.3722 | -104.4371
15.3759 | -104.4328
15.3796 | -104.4281
15.3832 | -104.4223
15.3869 | -104.4179
15.3905 | -104.4182
15.3942 | -104.4162
15.3978 | -104.4109
15.4015 | -104.408
15.4051 | -104.408
15.4088 | -104.4091
15.4124 | -104.4074
15.4161 | -104.4043
15.4198 | -104.4026
15.4234 | -104.3997
15.4271 | -104.3993
15.4307 | -104.4037
15.4344 | -104.4046
15.438 | -104.402
15.4417 | -104.3963
15.4453 | -104.3886
15.449 | -104.3839
15.4527 | -104.3826
15.4563 | -104.3796
15.46 | -104.3718
15.4636 | -104.3661
15.4673 | -104.3599
15.4709 | -104.3539
15.4746 | -104.3521
15.4782 | -104.351
15.4819 | -104.3488
15.4855 | -104.3449
15.4892 | -104.3422
15.4929 | -104.3413
15.4965 | -104.3449
15.5002 | -104.3489
15.5038 | -104.3518
15.5075 | -104.355
15.5111 | -104.203
15.5148 | -103.255
15.5184 | -102.7696
15.5221 | -102.5223
15.5257 | -102.256
15.5294 | -101.9712
15.5331 | -101.6691
15.5367 | -101.3534
15.5404 | -101.0273
15.544 | -100.693
15.5477 | -100.3539
15.5513 | -100.0132
15.555 | -99.6739
15.5586 | -99.3384
15.5623 | -99.0096
15.566 | -98.6891
15.5696 | -98.3791
15.5733 | -98.0815
15.5769 | -97.799
15.5806 | -97.5324
15.5842 | -97.2824
15.5879 | -97.0509
15.5915 | -96.8389
15.5952 | -96.6467
15.5988 | -96.4749
15.6025 | -96.3244
15.6062 | -96.1958
15.6098 | -96.089
15.6135 | -96.0043
15.6171 | -95.9427
15.6208 | -95.9044
15.6244 | -95.8884
15.6281 | -95.8952
15.6317 | -95.9252
15.6354 | -95.9779
15.639 | -96.0536
15.6427 | -96.1518
15.6464 | -96.2719
15.65 | -96.4136
15.6537 | -96.5766
15.6573 | -96.7605
15.661 | -96.9648
15.6646 | -97.1893
15.6683 | -97.432
15.6719 | -97.6917
15.6756 | -97.9676
15.6793 | -98.2588
15.6829 | -98.564
15.6866 | -98.8808
15.6902 | -99.2062
15.6939 | -99.5386
15.6975 | -99.8763
15.7012 | -100.2157
15.7048 | -100.5545
15.7085 | -100.8901
15.7121 | -101.2194
15.7158 | -101.54
15.7195 | -101.8468
15.7231 | -102.1378
15.7268 | -102.4113
15.7304 | -102.6653
15.7341 | -102.9951
15.7377 | -103.9364
15.7414 | -104.3388
15.745 | -104.3467
15.7487 | -104.3547
15.7523 | -104.3604
15.756 | -104.3682
15.7597 | -104.3767
15.7633 | -104.3845
15.767 | -104.3948
15.7706 | -104.4077
15.7743 | -104.4209
15.7779 | -104.4318
15.7816 | -104.444
15.7852 | -104.4553
15.7889 | -104.468
15.7926 | -104.4821
15.7962 | -104.496
15.7999 | -104.5122
15.8035 | -104.5265
15.8072 | -104.5423
15.8108 | -104.5568
15.8145 | -104.5702
15.8181 | -104.5827
15.8218 | -104.5976
15.8254 | -104.6138
15.8291 | -104.6265
15.8328 | -104.6356
15.8364 | -104.6464
15.8401 | -104.6591
15.8437 | -104.6702
15.8474 | -104.6769
15.851 | -104.6839
15.8547 | -104.6906
15.8583 | -104.697
15.862 | -104.7024
15.8656 | -104.7058
15.8693 | -104.7071
15.873 | -104.7075
15.8766 | -104.7039
15.8803 | -104.6973
15.8839 | -104.6889
15.8876 | -104.6738
15.8912 | -104.6619
15.8949 | -104.6514
15.8985 | -104.6416
15.9022 | -104.6293
15.9059 | -104.6158
15.9095 | -104.5979
15.9132 | -104.5785
15.9168 | -104.5624
15.9205 | -104.5462
15.9241 | -104.5269
15.9278 | -104.5092
15.9314 | -104.4887
15.9351 | -104.4675
15.9387 | -104.4472
15.9424 | -104.4256
15.9461 | -104.4045
15.9497 | -104.3825
15.9534 | -104.3623
15.957 | -104.3405
15.9607 | -104.3177
15.9643 | -104.2968
15.968 | -104.2738
15.9716 | -104.2516
15.9753 | -104.2276
15.9789 | -104.2063
15.9826 | -104.1863
15.9863 | -104.17
15.9899 | -104.1539
15.9936 | -104.1365
15.9972 | -104.1185
16.0009 | -104.1001
16.0045 | -104.0864
16.0082 | -104.0752
16.0118 | -104.0659
16.0155 | -104.0563
16.0192 | -104.0482
16.0228 | -104.0401
16.0265 | -104.0289
16.0301 | -104.0214
16.0338 | -104.0168
16.0374 | -104.0126
16.0411 | -104.01
16.0447 | -104.008
16.0484 | -104.0043
16.052 | -104.0002
16.0557 | -103.9963
16.0594 | -103.9934
16.063 | -103.9892
16.0667 | -103.9853
16.0703 | -103.9825
16.074 | -103.9801
16.0776 | -103.9808
16.0813 | -103.979
16.0849 | -103.9758
16.0886 | -103.9747
16.0922 | -103.975
16.0959 | -103.9761
16.0996 | -103.9787
16.1032 | -103.9835
16.1069 | -103.9887
16.1105 | -103.9968
16.1142 | -104.0067
16.1178 | -104.0187
16.1215 | -104.0274
16.1251 | -104.0365
16.1288 | -104.0456
16.1325 | -104.0584
16.1361 | -104.0725
16.1398 | -104.0885
16.1434 | -104.1054
16.1471 | -104.1239
16.1507 | -104.1412
16.1544 | -104.1607
16.158 | -104.1818
16.1617 | -104.2067
16.1653 | -104.2311
16.169 | -104.2549
16.1727 | -104.2792
16.1763 | -104.3038
16.18 | -104.3316
16.1836 | -104.3573
16.1873 | -104.3838
16.1909 | -104.4102
16.1946 | -104.4333
16.1982 | -104.4577
16.2019 | -104.4818
16.2055 | -104.5062
16.2092 | -104.5291
16.2129 | -104.5478
16.2165 | -104.564
16.2202 | -104.5786
16.2238 | -104.5919
16.2275 | -104.6051
16.2311 | -104.6168
16.2348 | -104.6248
16.2384 | -104.6266
16.2421 | -104.6278
16.2458 | -104.6282
16.2494 | -104.6254
16.2531 | -104.6227
16.2567 | -104.6204
16.2604 | -104.6157
16.264 | -104.6071
16.2677 | -104.5991
16.2713 | -104.5896
16.275 | -104.5797
16.2786 | -104.5676
16.2823 | -104.5555
16.286 | -104.5434
16.2896 | -104.5305
16.2933 | -104.5221
16.2969 | -104.5115
16.3006 | -104.4981
16.3042 | -104.4869
16.3079 | -104.4767
16.3115 | -104.4689
16.3152 | -104.4624
16.3188 | -104.4542
16.3225 | -104.4469
16.3262 | -104.4429
16.3298 | -104.4384
16.3335 | -104.4339
16.3371 | -104.4315
16.3408 | -104.4274
16.3444 | -104.4277
16.3481 | -104.4258
16.3517 | -104.4262
16.3554 | -104.4259
16.3591 | -104.4262
16.3627 | -104.4263
16.3664 | -104.4261
16.37 | -104.4262
16.3737 | -104.4286
16.3773 | -104.4321
16.381 | -104.4349
16.3846 | -104.4407
16.3883 | -104.4471
16.3919 | -104.4516
16.3956 | -104.454
16.3993 | -104.4564
16.4029 | -104.4599
16.4066 | -104.4656
16.4102 | -104.4738
16.4139 | -104.4816
16.4175 | -104.4926
16.4212 | -104.5046
16.4248 | -104.5157
16.4285 | -104.5289
16.4321 | -104.5419
16.4358 | -104.5558
16.4395 | -104.5695
16.4431 | -104.5826
16.4468 | -104.5973
16.4504 | -104.6108
16.4541 | -104.6242
16.4577 | -104.6389
16.4614 | -104.6526
16.465 | -104.6673
16.4687 | -104.6801
16.4724 | -104.6947
16.476 | -104.7056
16.4797 | -104.7163
16.4833 | -104.7262
16.487 | -104.735
16.4906 | -104.7412
16.4943 | -104.7463
16.4979 | -104.7539
16.5016 | -104.7564
16.5052 | -104.7585
16.5089 | -104.7596
16.5126 | -104.7601
16.5162 | -104.7618
16.5199 | -104.7611
16.5235 | -104.7576
16.5272 | -104.7493
16.5308 | -104.7415
16.5345 | -104.7344
16.5381 | -104.7271
16.5418 | -104.7191
16.5454 | -104.7111
16.5491 | -104.701
16.5528 | -104.6902
16.5564 | -104.6831
16.5601 | -104.6741
16.5637 | -104.6661
16.5674 | -104.6564
16.571 | -104.6465
16.5747 | -104.6391
16.5783 | -104.6316
16.582 | -104.624
16.5857 | -104.6168
16.5893 | -104.6098
16.593 | -104.6029
16.5966 | -104.5961
16.6003 | -104.5904
16.6039 | -104.5853
16.6076 | -104.5829
16.6112 | -104.5801
16.6149 | -104.5775
16.6185 | -104.575
16.6222 | -104.5731
16.6259 | -104.5719
16.6295 | -104.573
16.6332 | -104.5769
16.6368 | -104.5844
16.6405 | -104.5889
16.6441 | -104.5908
16.6478 | -104.5942
16.6514 | -104.5989
16.6551 | -104.6031
16.6587 | -104.6107
16.6624 | -104.6179
16.6661 | -104.6253
16.6697 | -104.6363
16.6734 | -104.6475
16.677 | -104.6584
16.6807 | -104.6684
16.6843 | -104.6801
16.688 | -104.6944
16.6916 | -104.7061
16.6953 | -104.7148
16.699 | -104.7236
16.7026 | -104.7313
16.7063 | -104.7388
16.7099 | -104.7453
16.7136 | -104.7513
16.7172 | -104.7549
16.7209 | -104.759
16.7245 | -104.7631
16.7282 | -104.7674
16.7318 | -104.771
16.7355 | -104.7727
16.7392 | -104.7745
16.7428 | -104.7756
16.7465 | -104.7759
16.7501 | -104.7798
16.7538 | -104.7819
16.7574 | -104.7837
16.7611 | -104.7888
16.7647 | -104.7956
16.7684 | -104.7996
16.772 | -104.803
16.7757 | -104.8069
16.7794 | -104.8121
16.783 | -104.8211
16.7867 | -104.831
16.7903 | -104.8428
16.794 | -104.8509
16.7976 | -104.8618
16.8013 | -104.8738
16.8049 | -104.8891
16.8086 | -104.9047
16.8123 | -104.9204
16.8159 | -104.9343
16.8196 | -104.9482
16.8232 | -104.9659
16.8269 | -104.983
16.8305 | -105.0003
16.8342 | -105.0174
16.8378 | -105.0348
16.8415 | -105.0524
16.8451 | -105.07
16.8488 | -105.0868
16.8525 | -105.1016
16.8561 | -105.1182
16.8598 | -105.1323
16.8634 | -105.1467
16.8671 | -105.1623
16.8707 | -105.1697
16.8744 | -105.181
16.878 | -105.1862
16.8817 | -105.1898
16.8853 | -105.1945
16.889 | -105.1962
16.8927 | -105.1986
16.8963 | -105.199
16.9 | -105.1954
16.9036 | -105.1905
16.9073 | -105.1826
16.9109 | -105.1737
16.9146 | -105.1626
16.9182 | -105.15
16.9219 | -105.1385
16.9256 | -105.1246
16.9292 | -105.1085
16.9329 | -105.0941
16.9365 | -105.0788
16.9402 | -105.0589
16.9438 | -105.0389
16.9475 | -105.0203
16.9511 | -105.0009
16.9548 | -104.9821
16.9584 | -104.9659
16.9621 | -104.9489
16.9658 | -104.9312
16.9694 | -104.9155
16.9731 | -104.8988
16.9767 | -104.8808
16.9804 | -104.8681
16.984 | -104.856
16.9877 | -104.8461
16.9913 | -104.8374
16.995 | -104.8265
16.9986 | -104.8132
17.0023 | -104.8042
17.006 | -104.798
17.0096 | -104.7933
17.0133 | -104.7912
17.0169 | -104.7873
17.0206 | -104.7849
17.0242 | -104.7812
17.0279 | -104.7806
17.0315 | -104.7798
17.0352 | -104.7798
17.0389 | -104.7804
17.0425 | -104.782
17.0462 | -104.7841
17.0498 | -104.7868
17.0535 | -104.7876
17.0571 | -104.7911
17.0608 | -104.7953
17.0644 | -104.7959
17.0681 | -104.7989
17.0717 | -104.8032
17.0754 | -104.8086
17.0791 | -104.813
17.0827 | -104.8161
17.0864 | -104.8191
17.09 | -104.8233
17.0937 | -104.8301
17.0973 | -104.8347
17.101 | -104.8396
17.1046 | -104.845
17.1083 | -104.8505
17.1119 | -104.8533
17.1156 | -104.8571
17.1193 | -104.8629
17.1229 | -104.8711
17.1266 | -104.8827
17.1302 | -104.892
17.1339 | -104.9002
17.1375 | -104.9066
17.1412 | -104.9145
17.1448 | -104.922
17.1485 | -104.9293
17.1522 | -104.9368
17.1558 | -104.9426
17.1595 | -104.9499
17.1631 | -104.9565
17.1668 | -104.9601
17.1704 | -104.9625
17.1741 | -104.9653
17.1777 | -104.969
17.1814 | -104.9735
17.185 | -104.9788
17.1887 | -104.9812
17.1924 | -104.9802
17.196 | -104.9793
17.1997 | -104.9795
17.2033 | -104.9782
17.207 | -104.9776
17.2106 | -104.9761
17.2143 | -104.9726
17.2179 | -104.9702
17.2216 | -104.968
17.2252 | -104.9643
17.2289 | -104.9622
17.2326 | -104.9591
17.2362 | -104.9535
17.2399 | -104.9508
17.2435 | -104.9519
17.2472 | -104.9472
17.2508 | -104.9418
17.2545 | -104.9378
17.2581 | -104.9343
17.2618 | -104.932
17.2655 | -104.9287
17.2691 | -104.9281
17.2728 | -104.9273
17.2764 | -104.9242
17.2801 | -104.9218
17.2837 | -104.9168
17.2874 | -104.9125
17.291 | -104.9104
17.2947 | -104.9108
17.2983 | -104.9122
17.302 | -104.9115
17.3057 | -104.9105
17.3093 | -104.9097
17.313 | -104.9092
17.3166 | -104.9097
17.3203 | -104.9104
17.3239 | -104.9125
17.3276 | -104.9133
17.3312 | -104.9137
17.3349 | -104.9132
17.3385 | -104.9107
17.3422 | -104.9068
17.3459 | -104.9055
17.3495 | -104.9039
17.3532 | -104.9016
17.3568 | -104.8991
17.3605 | -104.896
17.3641 | -104.8918
17.3678 | -104.8853
17.3714 | -104.876
17.3751 | -104.8666
17.3788 | -104.8554
17.3824 | -104.8459
17.3861 | -104.8353
17.3897 | -104.8265
17.3934 | -104.8196
17.397 | -104.8109
17.4007 | -104.8035
17.4043 | -104.7946
17.408 | -104.7852
17.4116 | -104.7785
17.4153 | -104.7745
17.419 | -104.7708
17.4226 | -104.764
17.4263 | -104.7582
17.4299 | -104.7541
17.4336 | -104.7507
17.4372 | -104.7488
17.4409 | -104.7482
17.4445 | -104.7489
17.4482 | -104.7468
17.4518 | -104.7454
17.4555 | -104.7467
17.4592 | -104.7494
17.4628 | -104.754
17.4665 | -104.7582
17.4701 | -104.7653
17.4738 | -104.7705
17.4774 | -104.7744
17.4811 | -104.7803
17.4847 | -104.7886
17.4884 | -104.7976
17.4921 | -104.8055
17.4957 | -104.8164
17.4994 | -104.8269
17.503 | -104.8342
17.5067 | -104.8388
17.5103 | -104.8486
17.514 | -104.8586
17.5176 | -104.869
17.5213 | -104.8797
17.5249 | -104.8925
17.5286 | -104.9049
17.5323 | -104.9148
17.5359 | -104.9254
17.5396 | -104.9353
17.5432 | -104.945
17.5469 | -104.9559
17.5505 | -104.9685
17.5542 | -104.9803
17.5578 | -104.9912
17.5615 | -105.0022
17.5651 | -105.0104
17.5688 | -105.0171
17.5725 | -105.0237
17.5761 | -105.0312
17.5798 | -105.0402
17.5834 | -105.0503
17.5871 | -105.0596
17.5907 | -105.0685
17.5944 | -105.0738
17.598 | -105.0785
17.6017 | -105.0814
17.6054 | -105.0871
17.609 | -105.0919
17.6127 | -105.0959
17.6163 | -105.0979
17.62 | -105.099
17.6236 | -105.1006
17.6273 | -105.1009
17.6309 | -105.0988
17.6346 | -105.0942
17.6382 | -105.0888
17.6419 | -105.08
17.6456 | -105.0726
17.6492 | -105.0702
17.6529 | -105.0663
17.6565 | -105.0622
17.6602 | -105.0556
17.6638 | -105.05
17.6675 | -105.0445
17.6711 | -105.039
17.6748 | -105.0354
17.6784 | -105.0322
17.6821 | -105.0275
17.6858 | -105.0231
17.6894 | -105.0183
17.6931 | -105.0147
17.6967 | -105.0112
17.7004 | -105.01
17.704 | -105.009
17.7077 | -105.0067
17.7113 | -105.0065
17.715 | -105.0054
17.7187 | -105.0056
17.7223 | -105.0055
17.726 | -105.0066
17.7296 | -105.0086
17.7333 | -105.0097
17.7369 | -105.0113
17.7406 | -105.0142
17.7442 | -105.0161
17.7479 | -105.0177
17.7515 | -105.0199
17.7552 | -105.0238
17.7589 | -105.0306
17.7625 | -105.0348
17.7662 | -105.0352
17.7698 | -105.0348
17.7735 | -105.0349
17.7771 | -105.0367
17.7808 | -105.0402
17.7844 | -105.0456
17.7881 | -105.0476
17.7917 | -105.0458
17.7954 | -105.0424
17.7991 | -105.0389
17.8027 | -105.0332
17.8064 | -105.0303
17.81 | -105.0286
17.8137 | -105.0234
17.8173 | -105.0163
17.821 | -105.0105
17.8246 | -105.0053
17.8283 | -105
17.832 | -104.9925
17.8356 | -104.9842
17.8393 | -104.9763
17.8429 | -104.9682
17.8466 | -104.9601
17.8502 | -104.9504
17.8539 | -104.9424
17.8575 | -104.9333
17.8612 | -104.9265
17.8648 | -104.9208
17.8685 | -104.9156
17.8722 | -104.9092
17.8758 | -104.894
17.8795 | -104.8857
17.8831 | -104.8814
17.8868 | -104.8791
17.8904 | -104.876
17.8941 | -104.8753
17.8977 | -104.8743
17.9014 | -104.8698
17.905 | -104.8639
17.9087 | -104.8584
17.9124 | -104.8563
17.916 | -104.8522
17.9197 | -104.8452
17.9233 | -104.8403
17.927 | -104.8335
17.9306 | -104.8306
17.9343 | -104.8267
17.9379 | -104.8224
17.9416 | -104.818
17.9453 | -104.8133
17.9489 | -104.8087
17.9526 | -104.8044
17.9562 | -104.7983
17.9599 | -104.7904
17.9635 | -104.7807
17.9672 | -104.7709
17.9708 | -104.7625
17.9745 | -104.7542
17.9781 | -104.7434
17.9818 | -104.7349
17.9855 | -104.7239
17.9891 | -104.7136
17.9928 | -104.7033
17.9964 | -104.6917
18.0001 | -104.6826
18.0037 | -104.6724
18.0074 | -104.6603
18.011 | -104.6506
18.0147 | -104.6384
18.0183 | -104.6244
18.022 | -104.6149
18.0257 | -104.604
18.0293 | -104.5905
18.033 | -104.5777
18.0366 | -104.5647
18.0403 | -104.5521
18.0439 | -104.5395
18.0476 | -104.5278
18.0512 | -104.5179
18.0549 | -104.5053
18.0586 | -104.4924
18.0622 | -104.4772
18.0659 | -104.4627
18.0695 | -104.4497
18.0732 | -104.4389
18.0768 | -104.4268
18.0805 | -104.4124
18.0841 | -104.3964
18.0878 | -104.3788
18.0914 | -104.3606
18.0951 | -104.3435
18.0988 | -104.3297
18.1024 | -104.3176
18.1061 | -104.3043
18.1097 | -104.2883
18.1134 | -104.2718
18.117 | -104.2556
18.1207 | -104.238
18.1243 | -104.2213
18.128 | -104.2054
18.1316 | -104.1927
18.1353 | -104.1739
18.139 | -104.1548
18.1426 | -104.1377
18.1463 | -104.1193
18.1499 | -104.0988
18.1536 | -104.0792
18.1572 | -104.0616
18.1609 | -104.0448
18.1645 | -104.0293
18.1682 | -104.0122
18.1718 | -103.9952
18.1755 | -103.9809
18.1792 | -103.9657
18.1828 | -103.9503
18.1865 | -103.9342
18.1901 | -103.9166
18.1938 | -103.9019
18.1974 | -103.8888
18.2011 | -103.8759
18.2047 | -103.863
18.2084 | -103.8495
18.2121 | -103.8358
18.2157 | -103.8253
18.2194 | -103.8167
18.223 | -103.8084
18.2267 | -103.8001
18.2303 | -103.7925
18.234 | -103.7828
18.2376 | -103.777
18.2413 | -103.7711
18.2449 | -103.7653
18.2486 | -103.7592
18.2523 | -103.7541
18.2559 | -103.7508
18.2596 | -103.7452
18.2632 | -103.7374
18.2669 | -103.7325
18.2705 | -103.7287
18.2742 | -103.7263
18.2778 | -103.7229
18.2815 | -103.7167
18.2852 | -103.7131
18.2888 | -103.7078
18.2925 | -103.7007
18.2961 | -103.6939
18.2998 | -103.6886
18.3034 | -103.6845
18.3071 | -103.6825
18.3107 | -103.6765
18.3144 | -103.6709
18.318 | -103.6685
18.3217 | -103.6662
18.3254 | -103.6615
18.329 | -103.6576
18.3327 | -103.6532
18.3363 | -103.6481
18.34 | -103.6428
18.3436 | -103.6374
18.3473 | -103.6326
18.3509 | -103.6283
18.3546 | -103.6214
18.3582 | -103.6126
18.3619 | -103.6039
18.3656 | -103.5977
18.3692 | -103.5941
18.3729 | -103.5892
18.3765 | -103.5806
18.3802 | -103.5743
18.3838 | -103.5699
18.3875 | -103.5665
18.3911 | -103.5643
18.3948 | -103.5618
18.3985 | -103.5568
18.4021 | -103.5513
18.4058 | -103.5465
18.4094 | -103.5447
18.4131 | -103.5449
18.4167 | -103.5451
18.4204 | -103.5465
18.424 | -103.5453
18.4277 | -103.5451
18.4313 | -103.545
18.435 | -103.5446
18.4387 | -103.5452
18.4423 | -103.5437
18.446 | -103.5436
18.4496 | -103.5438
18.4533 | -103.5399
18.4569 | -103.5381
18.4606 | -103.537
18.4642 | -103.5358
18.4679 | -103.5359
18.4715 | -103.5351
18.4752 | -103.5346
18.4789 | -103.5318
18.4825 | -103.5301
18.4862 | -103.5286
18.4898 | -103.5271
18.4935 | -103.5231
18.4971 | -103.5182
18.5008 | -103.5135
18.5044 | -103.5078
18.5081 | -103.502
18.5117 | -103.4989
18.5154 | -103.4962
18.5191 | -103.494
18.5227 | -103.4909
18.5264 | -103.4928
18.53 | -103.4908
18.5337 | -103.4905
18.5373 | -103.4892
18.541 | -103.4895
18.5446 | -103.4916
18.5483 | -103.4927
18.552 | -103.4964
18.5556 | -103.5006
18.5593 | -103.503
18.5629 | -103.5055
18.5666 | -103.5067
18.5702 | -103.5046
18.5739 | -103.5058
18.5775 | -103.5075
18.5812 | -103.5075
18.5848 | -103.5077
18.5885 | -103.5095
18.5922 | -103.5116
18.5958 | -103.5136
18.5995 | -103.516
18.6031 | -103.5162
18.6068 | -103.5176
18.6104 | -103.5179
18.6141 | -103.5151
18.6177 | -103.5104
18.6214 | -103.5059
18.6251 | -103.4987
18.6287 | -103.4918
18.6324 | -103.4852
18.636 | -103.4787
18.6397 | -103.474
18.6433 | -103.4657
18.647 | -103.4542
18.6506 | -103.4405
18.6543 | -103.4283
18.6579 | -103.4173
18.6616 | -103.4024
18.6653 | -103.3869
18.6689 | -103.3736
18.6726 | -103.3588
18.6762 | -103.3419
18.6799 | -103.3232
18.6835 | -103.3048
18.6872 | -103.2895
18.6908 | -103.2711
18.6945 | -103.2512
18.6981 | -103.2313
18.7018 | -103.2115
18.7055 | -103.195
18.7091 | -103.1772
18.7128 | -103.1577
18.7164 | -103.1393
18.7201 | -103.1196
18.7237 | -103.1008
18.7274 | -103.0822
18.731 | -103.0658
18.7347 | -103.0489
18.7384 | -103.0313
18.742 | -103.0156
18.7457 | -103.0016
18.7493 | -102.9866
18.753 | -102.9719
18.7566 | -102.9573
18.7603 | -102.9479
18.7639 | -102.9398
18.7676 | -102.9273
18.7712 | -102.919
18.7749 | -102.9142
18.7786 | -102.9065
18.7822 | -102.8989
18.7859 | -102.8914
18.7895 | -102.8847
18.7932 | -102.8779
18.7968 | -102.8746
18.8005 | -102.8706
18.8041 | -102.8676
18.8078 | -102.866
18.8114 | -102.8651
18.8151 | -102.8638
18.8188 | -102.8617
18.8224 | -102.8594
18.8261 | -102.8589
18.8297 | -102.8576
18.8334 | -102.8582
18.837 | -102.8589
18.8407 | -102.8564
18.8443 | -102.8552
18.848 | -102.855
18.8516 | -102.8542
18.8553 | -102.8537
18.859 | -102.8526
18.8626 | -102.8502
18.8663 | -102.8465
18.8699 | -102.8458
18.8736 | -102.8463
18.8772 | -102.8475
18.8809 | -102.8488
18.8845 | -102.8526
18.8882 | -102.8562
18.8919 | -102.8596
18.8955 | -102.8608
18.8992 | -102.8598
18.9028 | -102.8612
18.9065 | -102.8619
18.9101 | -102.8639
18.9138 | -102.8685
18.9174 | -102.873
18.9211 | -102.8765
18.9247 | -102.879
18.9284 | -102.8842
18.9321 | -102.889
18.9357 | -102.8916
18.9394 | -102.8956
18.943 | -102.8985
18.9467 | -102.9038
18.9503 | -102.9102
18.954 | -102.9156
18.9576 | -102.9239
18.9613 | -102.933
18.965 | -102.941
18.9686 | -102.9506
18.9723 | -102.9592
18.9759 | -102.9689
18.9796 | -102.9854
18.9832 | -102.9992
18.9869 | -103.012
18.9905 | -103.0265
18.9942 | -103.0421
18.9978 | -103.0591
19.0015 | -103.0737
19.0052 | -103.0877
19.0088 | -103.105
19.0125 | -103.1234
19.0161 | -103.1418
19.0198 | -103.1608
19.0234 | -103.1805
19.0271 | -103.2011
19.0307 | -103.2181
19.0344 | -103.2388
19.038 | -103.2606
19.0417 | -103.2775
19.0454 | -103.2929
19.049 | -103.3096
19.0527 | -103.3258
19.0563 | -103.3417
19.06 | -103.3601
19.0636 | -103.3751
19.0673 | -103.388
19.0709 | -103.3977
19.0746 | -103.4077
19.0783 | -103.4162
19.0819 | -103.4235
19.0856 | -103.4298
19.0892 | -103.4362
19.0929 | -103.441
19.0965 | -103.4435
19.1002 | -103.4444
19.1038 | -103.4439
19.1075 | -103.4428
19.1111 | -103.4431
19.1148 | -103.441
19.1185 | -103.437
19.1221 | -103.4342
19.1258 | -103.4295
19.1294 | -103.4235
19.1331 | -103.4182
19.1367 | -103.4134
19.1404 | -103.4088
19.144 | -103.4013
19.1477 | -103.3929
19.1513 | -103.3845
19.155 | -103.3777
19.1587 | -103.3683
19.1623 | -103.3598
19.166 | -103.3535
19.1696 | -103.3499
19.1733 | -103.3443
19.1769 | -103.3394
19.1806 | -103.3361
19.1842 | -103.3322
19.1879 | -103.3289
19.1915 | -103.3244
19.1952 | -103.3218
19.1989 | -103.3189
19.2025 | -103.3168
19.2062 | -103.3156
19.2098 | -103.3125
19.2135 | -103.31
19.2171 | -103.3062
19.2208 | -103.3033
19.2244 | -103.3025
19.2281 | -103.3025
19.2318 | -103.3025
19.2354 | -103.3025
19.2391 | -103.3014
19.2427 | -103.301
19.2464 | -103.2998
19.25 | -103.2996
19.2537 | -103.297
19.2573 | -103.2948
19.261 | -103.2949
19.2646 | -103.2926
19.2683 | -103.2902
19.272 | -103.2901
19.2756 | -103.2912
19.2793 | -103.2942
19.2829 | -103.2974
19.2866 | -103.2981
19.2902 | -103.2999
19.2939 | -103.3038
19.2975 | -103.3096
19.3012 | -103.3132
19.3048 | -103.3167
19.3085 | -103.32
19.3122 | -103.3252
19.3158 | -103.3311
19.3195 | -103.3367
19.3231 | -103.3417
19.3268 | -103.3452
19.3304 | -103.3481
19.3341 | -103.3511
19.3377 | -103.3554
19.3414 | -103.3601
19.3451 | -103.3648
19.3487 | -103.3674
19.3524 | -103.3698
19.356 | -103.3712
19.3597 | -103.371
19.3633 | -103.3705
19.367 | -103.3683
19.3706 | -103.3656
19.3743 | -103.3615
19.3779 | -103.3564
19.3816 | -103.3522
19.3853 | -103.3482
19.3889 | -103.3432
19.3926 | -103.3361
19.3962 | -103.3265
19.3999 | -103.3158
19.4035 | -103.306
19.4072 | -103.2969
19.4108 | -103.2869
19.4145 | -103.2739
19.4182 | -103.2582
19.4218 | -103.2431
19.4255 | -103.2296
19.4291 | -103.2104
19.4328 | -103.1917
19.4364 | -103.1758
19.4401 | -103.1591
19.4437 | -103.1409
19.4474 | -103.1231
19.451 | -103.1049
19.4547 | -103.088
19.4584 | -103.0709
19.462 | -103.0533
19.4657 | -103.0376
19.4693 | -103.0222
19.473 | -103.0088
19.4766 | -102.9965
19.4803 | -102.9859
19.4839 | -102.9786
19.4876 | -102.9735
19.4912 | -102.9688
19.4949 | -102.9636
19.4986 | -102.9613
19.5022 | -102.9589
19.5059 | -102.9751
19.5095 | -102.9781
19.5132 | -102.9823
19.5168 | -102.9893
19.5205 | -102.9975
19.5241 | -103.0073
19.5278 | -103.0207
19.5314 | -103.0377
19.5351 | -103.0547
19.5388 | -103.0718
19.5424 | -103.0903
19.5461 | -103.1077
19.5497 | -103.1241
19.5534 | -103.1429
19.557 | -103.1631
19.5607 | -103.1823
19.5643 | -103.2014
19.568 | -103.2205
19.5717 | -103.2382
19.5753 | -103.2553
19.579 | -103.2744
19.5826 | -103.2947
19.5863 | -103.3117
19.5899 | -103.3248
19.5936 | -103.3385
19.5972 | -103.3529
19.6009 | -103.366
19.6045 | -103.38
19.6082 | -103.3925
19.6119 | -103.4001
19.6155 | -103.4063
19.6192 | -103.4113
19.6228 | -103.4173
19.6265 | -103.4221
19.6301 | -103.4238
19.6338 | -103.4266
19.6374 | -103.4306
19.6411 | -103.4326
19.6447 | -103.4333
19.6484 | -103.433
19.6521 | -103.4326
19.6557 | -103.4298
19.6594 | -103.4256
19.663 | -103.4189
19.6667 | -103.4121
19.6703 | -103.4065
19.674 | -103.4004
19.6776 | -103.3952
19.6813 | -103.3893
19.685 | -103.3842
19.6886 | -103.3795
19.6923 | -103.3744
19.6959 | -103.3721
19.6996 | -103.3675
19.7032 | -103.3634
19.7069 | -103.3613
19.7105 | -103.3585
19.7142 | -103.3526
19.7178 | -103.3463
19.7215 | -103.3392
19.7252 | -103.3341
19.7288 | -103.3307
19.7325 | -103.3257
19.7361 | -103.3225
19.7398 | -103.3222
19.7434 | -103.3228
19.7471 | -103.3215
19.7507 | -103.3183
19.7544 | -103.3151
19.7581 | -103.3126
19.7617 | -103.3092
19.7654 | -103.3076
19.769 | -103.3075
19.7727 | -103.3081
19.7763 | -103.3064
19.78 | -103.3048
19.7836 | -103.3041
19.7873 | -103.2996
19.7909 | -103.2954
19.7946 | -103.2936
19.7983 | -103.2903
19.8019 | -103.286
19.8056 | -103.2816
19.8092 | -103.2757
19.8129 | -103.2691
19.8165 | -103.2633
19.8202 | -103.2564
19.8238 | -103.251
19.8275 | -103.2465
19.8311 | -103.2415
19.8348 | -103.2348
19.8385 | -103.2251
19.8421 | -103.2166
19.8458 | -103.2081
19.8494 | -103.199
19.8531 | -103.1898
19.8567 | -103.1789
19.8604 | -103.1648
19.864 | -103.1527
19.8677 | -103.1431
19.8713 | -103.1352
19.875 | -103.1263
19.8787 | -103.1155
19.8823 | -103.1049
19.886 | -103.0945
19.8896 | -103.0835
19.8933 | -103.0742
19.8969 | -103.0662
19.9006 | -103.0592
19.9042 | -103.0501
19.9079 | -103.0439
19.9116 | -103.0381
19.9152 | -103.029
19.9189 | -103.0171
19.9225 | -103.009
19.9262 | -103.0018
19.9298 | -102.9941
19.9335 | -102.986
19.9371 | -102.9776
19.9408 | -102.9709
19.9444 | -102.9651
19.9481 | -102.9562
19.9518 | -102.946
19.9554 | -102.9399
19.9591 | -102.9338
19.9627 | -102.9252
19.9664 | -102.9156
19.97 | -102.9067
19.9737 | -102.9012
19.9773 | -102.8961
19.981 | -102.8911
19.9846 | -102.8837
19.9883 | -102.8748
19.992 | -102.8676
19.9956 | -102.8611
19.9993 | -102.856
20.0029 | -102.8521
20.0066 | -102.8473
20.0102 | -102.8426
20.0139 | -102.8384
20.0175 | -102.836
20.0212 | -102.8342
20.0249 | -102.8295
20.0285 | -102.8258
20.0322 | -102.8257
20.0358 | -102.8295
20.0395 | -102.8308
20.0431 | -102.8305
20.0468 | -102.829
20.0504 | -102.8267
20.0541 | -102.8261
20.0577 | -102.8257
20.0614 | -102.8254
20.0651 | -102.8265
20.0687 | -102.8272
20.0724 | -102.8279
20.076 | -102.8282
20.0797 | -102.8283
20.0833 | -102.8276
20.087 | -102.8246
20.0906 | -102.8216
20.0943 | -102.8181
20.098 | -102.8131
20.1016 | -102.8086
20.1053 | -102.8064
20.1089 | -102.8023
20.1126 | -102.797
20.1162 | -102.7912
20.1199 | -102.7879
20.1235 | -102.7845
20.1272 | -102.7792
20.1308 | -102.7737
20.1345 | -102.7685
20.1382 | -102.7627
20.1418 | -102.757
20.1455 | -102.7523
20.1491 | -102.7451
20.1528 | -102.7373
20.1564 | -102.73
20.1601 | -102.7201
20.1637 | -102.7112
20.1674 | -102.7019
20.171 | -102.6934
20.1747 | -102.6861
20.1784 | -102.6784
20.182 | -102.6684
20.1857 | -102.6573
20.1893 | -102.6487
20.193 | -102.6412
20.1966 | -102.6355
20.2003 | -102.6299
20.2039 | -102.6236
20.2076 | -102.6182
20.2112 | -102.6124
20.2149 | -102.608
20.2186 | -102.6071
20.2222 | -102.6079
20.2259 | -102.6076
20.2295 | -102.6087
20.2332 | -102.6136
20.2368 | -102.6199
20.2405 | -102.6243
20.2441 | -102.6304
20.2478 | -102.6395
20.2515 | -102.6491
20.2551 | -102.6606
20.2588 | -102.6722
20.2624 | -102.6846
20.2661 | -102.6986
20.2697 | -102.7162
20.2734 | -102.7369
20.277 | -102.7557
20.2807 | -102.7735
20.2843 | -102.7912
20.288 | -102.8089
20.2917 | -102.8258
20.2953 | -102.8438
20.299 | -102.8622
20.3026 | -102.8783
20.3063 | -102.8955
20.3099 | -102.9158
20.3136 | -102.9354
20.3172 | -102.9521
20.3209 | -102.9682
20.3245 | -102.9852
20.3282 | -103.0022
20.3319 | -103.0181
20.3355 | -103.0346
20.3392 | -103.0491
20.3428 | -103.0617
20.3465 | -103.075
20.3501 | -103.089
20.3538 | -103.1024
20.3574 | -103.1125
20.3611 | -103.1228
20.3648 | -103.1348
20.3684 | -103.1451
20.3721 | -103.1546
20.3757 | -103.1646
20.3794 | -103.1719
20.383 | -103.1791
20.3867 | -103.1851
20.3903 | -103.1882
20.394 | -103.188
20.3976 | -103.1879
20.4013 | -103.1876
20.405 | -103.1879
20.4086 | -103.1854
20.4123 | -103.183
20.4159 | -103.1807
20.4196 | -103.1791
20.4232 | -103.176
20.4269 | -103.1733
20.4305 | -103.1715
20.4342 | -103.1701
20.4378 | -103.1672
20.4415 | -103.1621
20.4452 | -103.1565
20.4488 | -103.153
20.4525 | -103.1473
20.4561 | -103.1409
20.4598 | -103.1335
20.4634 | -103.1257
20.4671 | -103.1207
20.4707 | -103.1176
20.4744 | -103.1154
20.4781 | -103.1109
20.4817 | -103.1046
20.4854 | -103.1004
20.489 | -103.0992
20.4927 | -103.1002
20.4963 | -103.0987
20.5 | -103.0993
20.5036 | -103.1013
20.5073 | -103.1004
20.5109 | -103.0985
20.5146 | -103.098
20.5183 | -103.099
20.5219 | -103.0998
20.5256 | -103.0999
20.5292 | -103.1014
20.5329 | -103.1025
20.5365 | -103.1018
20.5402 | -103.1011
20.5438 | -103.101
20.5475 | -103.1
20.5511 | -103.0985
20.5548 | -103.0968
20.5585 | -103.0937
20.5621 | -103.0901
20.5658 | -103.0857
20.5694 | -103.0848
20.5731 | -103.0842
20.5767 | -103.0812
20.5804 | -103.0792
20.584 | -103.0789
20.5877 | -103.0733
20.5914 | -103.0684
20.595 | -103.066
20.5987 | -103.0626
20.6023 | -103.0578
20.606 | -103.055
20.6096 | -103.0532
20.6133 | -103.0492
20.6169 | -103.0417
20.6206 | -103.0346
20.6242 | -103.0303
20.6279 | -103.0264
20.6316 | -103.0227
20.6352 | -103.0173
20.6389 | -103.0112
20.6425 | -103.0037
20.6462 | -102.9972
20.6498 | -102.9929
20.6535 | -102.9862
20.6571 | -102.9774
20.6608 | -102.9705
20.6644 | -102.9629
20.6681 | -102.9562
20.6718 | -102.9493
20.6754 | -102.9401
20.6791 | -102.9322
20.6827 | -102.9253
20.6864 | -102.92
20.69 | -102.9147
20.6937 | -102.9085
20.6973 | -102.9004
20.701 | -102.8928
20.7047 | -102.885
20.7083 | -102.8782
20.712 | -102.8708
20.7156 | -102.8634
20.7193 | -102.8564
20.7229 | -102.85
20.7266 | -102.845
20.7302 | -102.8416
20.7339 | -102.8375
20.7375 | -102.8336
20.7412 | -102.8294
20.7449 | -102.8275
20.7485 | -102.8259
20.7522 | -102.8241
20.7558 | -102.8218
20.7595 | -102.8181
20.7631 | -102.8124
20.7668 | -102.8065
20.7704 | -102.8024
20.7741 | -102.8005
20.7777 | -102.7979
20.7814 | -102.793
20.7851 | -102.7882
20.7887 | -102.7824
20.7924 | -102.7763
20.796 | -102.7718
20.7997 | -102.7691
20.8033 | -102.7679
20.807 | -102.7655
20.8106 | -102.7608
20.8143 | -102.7569
20.818 | -102.7552
20.8216 | -102.7517
20.8253 | -102.7481
20.8289 | -102.7446
20.8326 | -102.7408
20.8362 | -102.7352
20.8399 | -102.7297
20.8435 | -102.7263
20.8472 | -102.7229
20.8508 | -102.7166
20.8545 | -102.7076
20.8582 | -102.7013
20.8618 | -102.6951
20.8655 | -102.6876
20.8691 | -102.6794
20.8728 | -102.6697
20.8764 | -102.6593
20.8801 | -102.6486
20.8837 | -102.6389
20.8874 | -102.6274
20.891 | -102.6138
20.8947 | -102.6009
20.8984 | -102.5897
20.902 | -102.5768
20.9057 | -102.5634
20.9093 | -102.5507
20.913 | -102.5396
20.9166 | -102.5273
20.9203 | -102.514
20.9239 | -102.5018
20.9276 | -102.4884
20.9313 | -102.4752
20.9349 | -102.4628
20.9386 | -102.4495
20.9422 | -102.4379
20.9459 | -102.4264
20.9495 | -102.4164
20.9532 | -102.4089
20.9568 | -102.3996
20.9605 | -102.3923
20.9641 | -102.3869
20.9678 | -102.3814
20.9715 | -102.3769
20.9751 | -102.3737
20.9788 | -102.3718
20.9824 | -102.3698
20.9861 | -102.3694
20.9897 | -102.3699
20.9934 | -102.3724
20.997 | -102.3742
21.0007 | -102.3771
21.0043 | -102.3823
21.008 | -102.389
21.0117 | -102.3977
21.0153 | -102.4059
21.019 | -102.4141
21.0226 | -102.4236
21.0263 | -102.4326
21.0299 | -102.4453
21.0336 | -102.4579
21.0372 | -102.4711
21.0409 | -102.4839
21.0446 | -102.4974
21.0482 | -102.5097
21.0519 | -102.5221
21.0555 | -102.5336
21.0592 | -102.5482
21.0628 | -102.5644
21.0665 | -102.5796
21.0701 | -102.5918
21.0738 | -102.6043
21.0774 | -102.6188
21.0811 | -102.6317
21.0848 | -102.6434
21.0884 | -102.6506
21.0921 | -102.662
21.0957 | -102.6709
21.0994 | -102.678
21.103 | -102.6848
21.1067 | -102.692
21.1103 | -102.6985
21.114 | -102.7057
21.1176 | -102.7122
21.1213 | -102.7187
21.125 | -102.7242
21.1286 | -102.7301
21.1323 | -102.7353
21.1359 | -102.739
21.1396 | -102.7421
21.1432 | -102.7415
21.1469 | -102.7414
21.1505 | -102.7453
21.1542 | -102.7475
21.1579 | -102.7453
21.1615 | -102.7437
21.1652 | -102.7415
21.1688 | -102.7409
21.1725 | -102.741
21.1761 | -102.7406
21.1798 | -102.7377
21.1834 | -102.7337
21.1871 | -102.7286
21.1907 | -102.7243
21.1944 | -102.7194
21.1981 | -102.7121
21.2017 | -102.7074
21.2054 | -102.702
21.209 | -102.6946
21.2127 | -102.6879
21.2163 | -102.6819
21.22 | -102.6747
21.2236 | -102.6655
21.2273 | -102.6573
21.2309 | -102.6501
21.2346 | -102.641
21.2383 | -102.6318
21.2419 | -102.6192
21.2456 | -102.6066
21.2492 | -102.5944
21.2529 | -102.5792
21.2565 | -102.5638
21.2602 | -102.5483
21.2638 | -102.5334
21.2675 | -102.5184
21.2712 | -102.5017
21.2748 | -102.4841
21.2785 | -102.4676
21.2821 | -102.4508
21.2858 | -102.4322
21.2894 | -102.414
21.2931 | -102.3967
21.2967 | -102.382
21.3004 | -102.3687
21.304 | -102.3536
21.3077 | -102.3378
21.3114 | -102.3222
21.315 | -102.3097
21.3187 | -102.2982
21.3223 | -102.2846
21.326 | -102.2735
21.3296 | -102.2637
21.3333 | -102.2556
21.3369 | -102.2498
21.3406 | -102.2457
21.3442 | -102.2429
21.3479 | -102.2423
21.3516 | -102.2421
21.3552 | -102.244
21.3589 | -102.2455
21.3625 | -102.2449
21.3662 | -102.247
21.3698 | -102.2525
21.3735 | -102.2574
21.3771 | -102.2636
21.3808 | -102.2706
21.3845 | -102.2791
21.3881 | -102.2853
21.3918 | -102.294
21.3954 | -102.3039
21.3991 | -102.3154
21.4027 | -102.3264
21.4064 | -102.3343
21.41 | -102.3448
21.4137 | -102.3555
21.4173 | -102.3634
21.421 | -102.3718
21.4247 | -102.381
21.4283 | -102.3907
21.432 | -102.3993
21.4356 | -102.4074
21.4393 | -102.4162
21.4429 | -102.423
21.4466 | -102.4306
21.4502 | -102.4364
21.4539 | -102.4401
21.4575 | -102.4432
21.4612 | -102.4457
21.4649 | -102.4462
21.4685 | -102.4459
21.4722 | -102.4447
21.4758 | -102.4432
21.4795 | -102.44
21.4831 | -102.4372
21.4868 | -102.4348
21.4904 | -102.4322
21.4941 | -102.4258
21.4978 | -102.4173
21.5014 | -102.4097
21.5051 | -102.3999
21.5087 | -102.3912
21.5124 | -102.3827
21.516 | -102.3721
21.5197 | -102.3623
21.5233 | -102.3528
21.527 | -102.3427
21.5306 | -102.3303
21.5343 | -102.3145
21.538 | -102.3014
21.5416 | -102.2896
21.5453 | -102.2771
21.5489 | -102.2644
21.5526 | -102.2528
21.5562 | -102.2407
21.5599 | -102.2291
21.5635 | -102.2148
21.5672 | -102.2011
21.5708 | -102.1888
21.5745 | -102.1757
21.5782 | -102.1618
21.5818 | -102.1468
21.5855 | -102.134
21.5891 | -102.1225
21.5928 | -102.1102
21.5964 | -102.0956
21.6001 | -102.0807
21.6037 | -102.0663
21.6074 | -102.0522
21.6111 | -102.039
21.6147 | -102.0262
21.6184 | -102.0145
21.622 | -102.0005
21.6257 | -101.9873
21.6293 | -101.976
21.633 | -101.9637
21.6366 | -101.9518
21.6403 | -101.9407
21.6439 | -101.928
21.6476 | -101.9166
21.6513 | -101.9076
21.6549 | -101.9007
21.6586 | -101.8962
21.6622 | -101.8924
21.6659 | -101.8899
21.6695 | -101.8889
21.6732 | -101.8882
21.6768 | -101.8879
21.6805 | -101.8902
21.6841 | -101.8943
21.6878 | -101.9
21.6915 | -101.9053
21.6951 | -101.9129
21.6988 | -101.9205
21.7024 | -101.9278
21.7061 | -101.9341
21.7097 | -101.9408
21.7134 | -101.9524
21.717 | -101.9639
21.7207 | -101.9769
21.7244 | -101.992
21.728 | -102.0069
21.7317 | -102.0201
21.7353 | -102.0346
21.739 | -102.0499
21.7426 | -102.0637
21.7463 | -102.0776
21.7499 | -102.0925
21.7536 | -102.1065
21.7572 | -102.1196
21.7609 | -102.1323
21.7646 | -102.1445
21.7682 | -102.1567
21.7719 | -102.1677
21.7755 | -102.1781
21.7792 | -102.1864
21.7828 | -102.1928
21.7865 | -102.1979
21.7901 | -102.2
21.7938 | -102.2013
21.7974 | -102.202
21.8011 | -102.2037
21.8048 | -102.205
21.8084 | -102.2038
21.8121 | -102.2021
21.8157 | -102.1996
21.8194 | -102.1963
21.823 | -102.1917
21.8267 | -102.1843
21.8303 | -102.1768
21.834 | -102.1693
21.8377 | -102.1629
21.8413 | -102.1568
21.845 | -102.1433
21.8486 | -102.1288
21.8523 | -102.1173
21.8559 | -102.1054
21.8596 | -102.0918
21.8632 | -102.0779
21.8669 | -102.0637
21.8705 | -102.0506
21.8742 | -102.0366
21.8779 | -102.0228
21.8815 | -102.0099
21.8852 | -101.9972
21.8888 | -101.9817
21.8925 | -101.9665
21.8961 | -101.9532
21.8998 | -101.9396
21.9034 | -101.9286
21.9071 | -101.9176
21.9107 | -101.9039
21.9144 | -101.8892
21.9181 | -101.875
21.9217 | -101.8616
21.9254 | -101.8481
21.929 | -101.8327
21.9327 | -101.8207
21.9363 | -101.8095
21.94 | -101.7993
21.9436 | -101.7904
21.9473 | -101.7794
21.951 | -101.769
21.9546 | -101.7613
21.9583 | -101.753
21.9619 | -101.745
21.9656 | -101.7363
21.9692 | -101.727
21.9729 | -101.7186
21.9765 | -101.7121
21.9802 | -101.7084
21.9838 | -101.7038
21.9875 | -101.6982
21.9912 | -101.6931
21.9948 | -101.6888
21.9985 | -101.6864
22.0021 | -101.6857
22.0058 | -101.6839
22.0094 | -101.6805
22.0131 | -101.6769
22.0167 | -101.6766
22.0204 | -101.6774
22.024 | -101.6772
22.0277 | -101.6743
22.0314 | -101.6677
22.035 | -101.6573
22.0387 | -101.65
22.0423 | -101.6461
22.046 | -101.6428
22.0496 | -101.6394
22.0533 | -101.6356
22.0569 | -101.6317
22.0606 | -101.6251
22.0643 | -101.6179
22.0679 | -101.6126
22.0716 | -101.6061
22.0752 | -101.5981
22.0789 | -101.5895
22.0825 | -101.5793
22.0862 | -101.568
22.0898 | -101.5585
22.0935 | -101.5491
22.0971 | -101.5401
22.1008 | -101.5294
22.1045 | -101.5175
22.1081 | -101.5051
22.1118 | -101.4918
22.1154 | -101.4778
22.1191 | -101.4648
22.1227 | -101.4511
22.1264 | -101.4348
22.13 | -101.4181
22.1337 | -101.403
22.1373 | -101.3862
22.141 | -101.3688
22.1447 | -101.3536
22.1483 | -101.3389
22.152 | -101.3248
22.1556 | -101.3099
22.1593 | -101.2979
22.1629 | -101.2869
22.1666 | -101.2739
22.1702 | -101.2591
22.1739 | -101.2451
22.1776 | -101.2336
22.1812 | -101.2219
22.1849 | -101.2102
22.1885 | -101.1988
22.1922 | -101.1872
22.1958 | -101.176
22.1995 | -101.1644
22.2031 | -101.1518
22.2068 | -101.1405
22.2104 | -101.1311
22.2141 | -101.1226
22.2178 | -101.1149
22.2214 | -101.1078
22.2251 | -101.1038
22.2287 | -101.0991
22.2324 | -101.093
22.236 | -101.087
22.2397 | -101.083
22.2433 | -101.0787
22.247 | -101.0706
22.2506 | -101.0623
22.2543 | -101.0551
22.258 | -101.0459
22.2616 | -101.0369
22.2653 | -101.0289
22.2689 | -101.02
22.2726 | -101.0122
22.2762 | -101.0029
22.2799 | -100.9928
22.2835 | -100.9838
22.2872 | -100.9741
22.2909 | -100.962
22.2945 | -100.9484
22.2982 | -100.9364
22.3018 | -100.9271
22.3055 | -100.9167
22.3091 | -100.9041
22.3128 | -100.8899
22.3164 | -100.8755
22.3201 | -100.8625
22.3237 | -100.8484
22.3274 | -100.8344
22.3311 | -100.8181
22.3347 | -100.8005
22.3384 | -100.7817
22.342 | -100.7641
22.3457 | -100.7485
22.3493 | -100.7292
22.353 | -100.7037
22.3566 | -100.6809
22.3603 | -100.6584
22.3639 | -100.6336
22.3676 | -100.6075
22.3713 | -100.5823
22.3749 | -100.5604
22.3786 | -100.5387
22.3822 | -100.5157
22.3859 | -100.4933
22.3895 | -100.4721
22.3932 | -100.451
22.3968 | -100.4314
22.4005 | -100.4111
22.4042 | -100.3885
22.4078 | -100.3668
22.4115 | -100.3448
22.4151 | -100.3262
22.4188 | -100.3082
22.4224 | -100.2901
22.4261 | -100.273
22.4297 | -100.2564
22.4334 | -100.2396
22.437 | -100.2223
22.4407 | -100.2066
22.4444 | -100.1934
22.448 | -100.1813
22.4517 | -100.1712
22.4553 | -100.1616
22.459 | -100.1502
22.4626 | -100.1385
22.4663 | -100.1282
22.4699 | -100.1191
22.4736 | -100.112
22.4772 | -100.1061
22.4809 | -100.0992
22.4846 | -100.0922
22.4882 | -100.0872
22.4919 | -100.0846
22.4955 | -100.0819
22.4992 | -100.0802
22.5028 | -100.0792
22.5065 | -100.0781
22.5101 | -100.0795
22.5138 | -100.0823
22.5175 | -100.0856
22.5211 | -100.0915
22.5248 | -100.0967
22.5284 | -100.1011
22.5321 | -100.1079
22.5357 | -100.1138
22.5394 | -100.121
22.543 | -100.1268
22.5467 | -100.1335
22.5503 | -100.1444
22.554 | -100.1565
22.5577 | -100.1679
22.5613 | -100.1808
22.565 | -100.1952
22.5686 | -100.2098
22.5723 | -100.2221
22.5759 | -100.2359
22.5796 | -100.2516
22.5832 | -100.2648
22.5869 | -100.2773
22.5905 | -100.2871
22.5942 | -100.2975
22.5979 | -100.3103
22.6015 | -100.321
22.6052 | -100.3315
22.6088 | -100.3429
22.6125 | -100.3532
22.6161 | -100.3616
22.6198 | -100.3685
22.6234 | -100.3755
22.6271 | -100.3839
22.6308 | -100.3931
22.6344 | -100.4005
22.6381 | -100.4059
22.6417 | -100.4125
22.6454 | -100.4181
22.649 | -100.4237
22.6527 | -100.4299
22.6563 | -100.4374
22.66 | -100.4454
22.6636 | -100.4525
22.6673 | -100.4629
22.671 | -100.4695
22.6746 | -100.481
22.6783 | -100.4879
22.6819 | -100.4941
22.6856 | -100.502
22.6892 | -100.5117
22.6929 | -100.5198
22.6965 | -100.526
22.7002 | -100.5307
22.7038 | -100.5368
22.7075 | -100.5418
22.7112 | -100.5475
22.7148 | -100.5542
22.7185 | -100.5608
22.7221 | -100.5682
22.7258 | -100.5768
22.7294 | -100.5852
22.7331 | -100.5932
22.7367 | -100.5992
22.7404 | -100.606
22.7441 | -100.6134
22.7477 | -100.6205
22.7514 | -100.6273
22.755 | -100.635
22.7587 | -100.6438
22.7623 | -100.6527
22.766 | -100.6603
22.7696 | -100.6652
22.7733 | -100.6708
22.7769 | -100.6779
22.7806 | -100.6862
22.7843 | -100.6956
22.7879 | -100.7051
22.7916 | -100.7142
22.7952 | -100.7234
22.7989 | -100.7321
22.8025 | -100.741
22.8062 | -100.7509
22.8098 | -100.7613
22.8135 | -100.77
22.8171 | -100.7783
22.8208 | -100.7866
22.8245 | -100.7968
22.8281 | -100.809
22.8318 | -100.8194
22.8354 | -100.8298
22.8391 | -100.841
22.8427 | -100.8536
22.8464 | -100.8645
22.85 | -100.8744
22.8537 | -100.8852
22.8574 | -100.8968
22.861 | -100.9074
22.8647 | -100.921
22.8683 | -100.9351
22.872 | -100.9453
22.8756 | -100.9559
22.8793 | -100.9667
22.8829 | -100.9773
22.8866 | -100.9882
22.8902 | -100.9992
22.8939 | -101.0108
22.8976 | -101.0224
22.9012 | -101.0332
22.9049 | -101.044
22.9085 | -101.0545
22.9122 | -101.0668
22.9158 | -101.0799
22.9195 | -101.0903
22.9231 | -101.1011
22.9268 | -101.1115
22.9304 | -101.1201
22.9341 | -101.1301
22.9378 | -101.1392
22.9414 | -101.1465
22.9451 | -101.1537
22.9487 | -101.159
22.9524 | -101.1638
22.956 | -101.168
22.9597 | -101.1722
22.9633 | -101.1771
22.967 | -101.1814
22.9707 | -101.1829
22.9743 | -101.1835
22.978 | -101.1828
22.9816 | -101.1814
22.9853 | -101.1795
22.9889 | -101.177
22.9926 | -101.1725
22.9962 | -101.1709
22.9999 | -101.169
23.0035 | -101.1671
23.0072 | -101.1634
23.0109 | -101.1597
23.0145 | -101.1557
23.0182 | -101.1511
23.0218 | -101.1475
23.0255 | -101.1446
23.0291 | -101.14
23.0328 | -101.1358
23.0364 | -101.1295
23.0401 | -101.1192
23.0437 | -101.109
23.0474 | -101.0979
23.0511 | -101.0882
23.0547 | -101.0802
23.0584 | -101.0722
23.062 | -101.0642
23.0657 | -101.0546
23.0693 | -101.0402
23.073 | -101.0242
23.0766 | -101.0116
23.0803 | -100.9997
23.084 | -100.9872
23.0876 | -100.9753
23.0913 | -100.9625
23.0949 | -100.9478
23.0986 | -100.9316
23.1022 | -100.9138
23.1059 | -100.8967
23.1095 | -100.88
23.1132 | -100.862
23.1168 | -100.844
23.1205 | -100.8266
23.1242 | -100.8096
23.1278 | -100.7932
23.1315 | -100.777
23.1351 | -100.7608
23.1388 | -100.7439
23.1424 | -100.7286
23.1461 | -100.7132
23.1497 | -100.6969
23.1534 | -100.6789
23.157 | -100.6638
23.1607 | -100.6503
23.1644 | -100.636
23.168 | -100.6194
23.1717 | -100.6059
23.1753 | -100.5958
23.179 | -100.5847
23.1826 | -100.5739
23.1863 | -100.5664
23.1899 | -100.5585
23.1936 | -100.5535
23.1973 | -100.5572
23.2009 | -100.5533
23.2046 | -100.5477
23.2082 | -100.5427
23.2119 | -100.5378
23.2155 | -100.5336
23.2192 | -100.5287
23.2228 | -100.525
23.2265 | -100.5235
23.2301 | -100.5221
23.2338 | -100.5222
23.2375 | -100.5231
23.2411 | -100.5241
23.2448 | -100.5236
23.2484 | -100.5236
23.2521 | -100.5253
23.2557 | -100.5269
23.2594 | -100.5295
23.263 | -100.5312
23.2667 | -100.5334
23.2703 | -100.5336
23.274 | -100.5326
23.2777 | -100.5313
23.2813 | -100.5291
23.285 | -100.5275
23.2886 | -100.5264
23.2923 | -100.5235
23.2959 | -100.5198
23.2996 | -100.5174
23.3032 | -100.5158
23.3069 | -100.511
23.3106 | -100.5049
23.3142 | -100.5
23.3179 | -100.4969
23.3215 | -100.4942
23.3252 | -100.4873
23.3288 | -100.479
23.3325 | -100.4693
23.3361 | -100.4623
23.3398 | -100.4565
23.3434 | -100.4485
23.3471 | -100.4419
23.3508 | -100.4341
23.3544 | -100.4262
23.3581 | -100.4184
23.3617 | -100.4116
23.3654 | -100.4071
23.369 | -100.4032
23.3727 | -100.3979
23.3763 | -100.3927
23.38 | -100.3867
23.3836 | -100.3807
23.3873 | -100.3765
23.391 | -100.3733
23.3946 | -100.3708
23.3983 | -100.3717
23.4019 | -100.3737
23.4056 | -100.3725
23.4092 | -100.3694
23.4129 | -100.3675
23.4165 | -100.3652
23.4202 | -100.3645
23.4239 | -100.3667
23.4275 | -100.3674
23.4312 | -100.3683
23.4348 | -100.3676
23.4385 | -100.3667
23.4421 | -100.3664
23.4458 | -100.3651
23.4494 | -100.3625
23.4531 | -100.3609
23.4567 | -100.36
23.4604 | -100.3607
23.4641 | -100.3604
23.4677 | -100.3614
23.4714 | -100.3624
23.475 | -100.3624
23.4787 | -100.3642
23.4823 | -100.3681
23.486 | -100.3691
23.4896 | -100.3681
23.4933 | -100.3716
23.4969 | -100.3768
23.5006 | -100.3806
23.5043 | -100.3835
23.5079 | -100.3871
23.5116 | -100.3928
23.5152 | -100.3959
23.5189 | -100.3985
23.5225 | -100.4013
23.5262 | -100.4047
23.5298 | -100.4118
23.5335 | -100.4188
23.5372 | -100.4273
23.5408 | -100.4376
23.5445 | -100.4496
23.5481 | -100.4611
23.5518 | -100.473
23.5554 | -100.4852
23.5591 | -100.4959
23.5627 | -100.5063
23.5664 | -100.5178
23.57 | -100.5298
23.5737 | -100.5415
23.5774 | -100.5534
23.581 | -100.5669
23.5847 | -100.5797
23.5883 | -100.5919
23.592 | -100.6016
23.5956 | -100.6112
23.5993 | -100.6191
23.6029 | -100.6273
23.6066 | -100.6347
23.6102 | -100.6434
23.6139 | -100.6532
23.6176 | -100.6632
23.6212 | -100.6707
23.6249 | -100.6745
23.6285 | -100.6779
23.6322 | -100.6826
23.6358 | -100.6857
23.6395 | -100.6904
23.6431 | -100.6935
23.6468 | -100.6952
23.6505 | -100.6937
23.6541 | -100.6922
23.6578 | -100.6919
23.6614 | -100.6904
23.6651 | -100.6884
23.6687 | -100.684
23.6724 | -100.6756
23.676 | -100.6681
23.6797 | -100.6588
23.6833 | -100.6498
23.687 | -100.6422
23.6907 | -100.6351
23.6943 | -100.6266
23.698 | -100.6169
23.7016 | -100.6075
23.7053 | -100.5976
23.7089 | -100.587
23.7126 | -100.577
23.7162 | -100.5662
23.7199 | -100.5546
23.7235 | -100.5431
23.7272 | -100.54
23.7309 | -100.5304
23.7345 | -100.5212
23.7382 | -100.5126
23.7418 | -100.5068
23.7455 | -100.503
23.7491 | -100.4997
23.7528 | -100.4982
23.7564 | -100.4968
23.7601 | -100.4959
23.7638 | -100.4983
23.7674 | -100.5018
23.7711 | -100.5073
23.7747 | -100.5138
23.7784 | -100.5212
23.782 | -100.5309
23.7857 | -100.5417
23.7893 | -100.5515
23.793 | -100.5614
23.7966 | -100.571
23.8003 | -100.5849
23.804 | -100.5995
23.8076 | -100.6147
23.8113 | -100.6326
23.8149 | -100.6504
23.8186 | -100.6691
23.8222 | -100.6902
23.8259 | -100.7135
23.8295 | -100.7351
23.8332 | -100.7557
23.8368 | -100.7781
23.8405 | -100.7996
23.8442 | -100.8215
23.8478 | -100.844
23.8515 | -100.8674
23.8551 | -100.8935
23.8588 | -100.9199
23.8624 | -100.9464
23.8661 | -100.9709
23.8697 | -100.9937
23.8734 | -101.0159
23.8771 | -101.0376
23.8807 | -101.0577
23.8844 | -101.0794
23.888 | -101.1007
23.8917 | -101.1201
23.8953 | -101.1371
23.899 | -101.1538
23.9026 | -101.1692
23.9063 | -101.1843
23.9099 | -101.1996
23.9136 | -101.2122
23.9173 | -101.222
23.9209 | -101.2318
23.9246 | -101.2408
23.9282 | -101.2509
23.9319 | -101.2594
23.9355 | -101.267
23.9392 | -101.2732
23.9428 | -101.2765
23.9465 | -101.2793
23.9501 | -101.2784
23.9538 | -101.2743
23.9575 | -101.272
23.9611 | -101.2713
23.9648 | -101.2689
23.9684 | -101.2629
23.9721 | -101.2562
23.9757 | -101.2491
23.9794 | -101.2427
23.983 | -101.2351
23.9867 | -101.227
23.9904 | -101.2166
23.994 | -101.2068
23.9977 | -101.198
24.0013 | -101.1888
24.005 | -101.1802
24.0086 | -101.1748
24.0123 | -101.1687
24.0159 | -101.1628
24.0196 | -101.1574
24.0232 | -101.1493
24.0269 | -101.1408
24.0306 | -101.1324
24.0342 | -101.1233
24.0379 | -101.115
24.0415 | -101.1078
24.0452 | -101.1031
24.0488 | -101.0985
24.0525 | -101.0925
24.0561 | -101.0875
24.0598 | -101.0847
24.0634 | -101.0809
24.0671 | -101.0763
24.0708 | -101.0746
24.0744 | -101.0713
24.0781 | -101.069
24.0817 | -101.0682
24.0854 | -101.0678
24.089 | -101.0667
24.0927 | -101.0651
24.0963 | -101.0645
24.1 | -101.0651
24.1037 | -101.0646
24.1073 | -101.0652
24.111 | -101.0666
24.1146 | -101.0687
24.1183 | -101.0712
24.1219 | -101.0736
24.1256 | -101.0768
24.1292 | -101.0807
24.1329 | -101.0855
24.1365 | -101.0912
24.1402 | -101.0938
24.1439 | -101.0958
24.1475 | -101.0988
24.1512 | -101.0997
24.1548 | -101.1
24.1585 | -101.1005
24.1621 | -101.1014
24.1658 | -101.1021
24.1694 | -101.1021
24.1731 | -101.1025
24.1767 | -101.1018
24.1804 | -101.1002
24.1841 | -101.0989
24.1877 | -101.0982
24.1914 | -101.0981
24.195 | -101.0966
24.1987 | -101.0944
24.2023 | -101.0914
24.206 | -101.09
24.2096 | -101.0903
24.2133 | -101.0904
24.217 | -101.0899
24.2206 | -101.0896
24.2243 | -101.0888
24.2279 | -101.0889
24.2316 | -101.0878
24.2352 | -101.0879
24.2389 | -101.0861
24.2425 | -101.0852
24.2462 | -101.0849
24.2498 | -101.0847
24.2535 | -101.0884
24.2572 | -101.0922
24.2608 | -101.0964
24.2645 | -101.0984
24.2681 | -101.1008
24.2718 | -101.1027
24.2754 | -101.1061
24.2791 | -101.1089
24.2827 | -101.1121
24.2864 | -101.1151
24.29 | -101.116
24.2937 | -101.1181
24.2974 | -101.1223
24.301 | -101.1239
24.3047 | -101.1235
24.3083 | -101.1175
24.312 | -101.1107
24.3156 | -101.106
24.3193 | -101.1024
24.3229 | -101.0985
24.3266 | -101.0946
24.3303 | -101.0921
24.3339 | -101.0905
24.3376 | -101.0877
24.3412 | -101.083
24.3449 | -101.0768
24.3485 | -101.071
24.3522 | -101.0644
24.3558 | -101.057
24.3595 | -101.0496
24.3631 | -101.0402
24.3668 | -101.0306
24.3705 | -101.0205
24.3741 | -101.0124
24.3778 | -101.0055
24.3814 | -100.9964
24.3851 | -100.987
24.3887 | -100.9757
24.3924 | -100.9626
24.396 | -100.9495
24.3997 | -100.9349
24.4033 | -100.9192
24.407 | -100.9042
24.4107 | -100.8887
24.4143 | -100.8723
24.418 | -100.8595
24.4216 | -100.8454
24.4253 | -100.8283
24.4289 | -100.8102
24.4326 | -100.7943
24.4362 | -100.7796
24.4399 | -100.7654
24.4436 | -100.7509
24.4472 | -100.7382
24.4509 | -100.7248
24.4545 | -100.7122
24.4582 | -100.7002
24.4618 | -100.6885
24.4655 | -100.6782
24.4691 | -100.6687
24.4728 | -100.6599
24.4764 | -100.651
24.4801 | -100.6422
24.4838 | -100.637
24.4874 | -100.6332
24.4911 | -100.6294
24.4947 | -100.6289
24.4984 | -100.6304
24.502 | -100.6324
24.5057 | -100.6345
24.5093 | -100.6381
24.513 | -100.6433
24.5166 | -100.6517
24.5203 | -100.6598
24.524 | -100.6693
24.5276 | -100.6791
24.5313 | -100.6892
24.5349 | -100.7017
24.5386 | -100.7145
24.5422 | -100.7269
24.5459 | -100.7419
24.5495 | -100.7584
24.5532 | -100.7744
24.5569 | -100.7908
24.5605 | -100.8076
24.5642 | -100.8242
24.5678 | -100.8411
24.5715 | -100.8555
24.5751 | -100.8668
24.5788 | -100.875
24.5824 | -100.8854
24.5861 | -100.8862
24.5897 | -100.8747
24.5934 | -100.8593
24.5971 | -100.8494
24.6007 | -100.8462
24.6044 | -100.8418
24.608 | -100.8381
24.6117 | -100.8332
24.6153 | -100.8247
24.619 | -100.8127
24.6226 | -100.7988
24.6263 | -100.7837
24.6299 | -100.7656
24.6336 | -100.748
24.6373 | -100.7274
24.6409 | -100.703
24.6446 | -100.6751
24.6482 | -100.6466
24.6519 | -100.6179
24.6555 | -100.5904
24.6592 | -100.5609
24.6628 | -100.5282
24.6665 | -100.4918
24.6702 | -100.4531
24.6738 | -100.4143
24.6775 | -100.3741
24.6811 | -100.3312
24.6848 | -100.2869
24.6884 | -100.2448
24.6921 | -100.2041
24.6957 | -100.1618
24.6994 | -100.1206
24.703 | -100.0814
24.7067 | -100.0422
24.7104 | -100.0054
24.714 | -99.9713
24.7177 | -99.9389
24.7213 | -99.9076
24.725 | -99.8781
24.7286 | -99.8533
24.7323 | -99.8301
24.7359 | -99.8099
24.7396 | -99.7923
24.7432 | -99.7782
24.7469 | -99.769
24.7506 | -99.7633
24.7542 | -99.762
24.7579 | -99.7648
24.7615 | -99.7706
24.7652 | -99.778
24.7688 | -99.7893
24.7725 | -99.8053
24.7761 | -99.8255
24.7798 | -99.879
24.7835 | -99.9047
24.7871 | -99.9321
24.7908 | -99.9633
24.7944 | -99.999
24.7981 | -100.0374
24.8017 | -100.0788
24.8054 | -100.1229
24.809 | -100.1694
24.8127 | -100.2195
24.8163 | -100.2721
24.82 | -100.3263
24.8237 | -100.3808
24.8273 | -100.4369
24.831 | -100.4947
24.8346 | -100.5514
24.8383 | -100.6101
24.8419 | -100.6693
24.8456 | -100.7281
24.8492 | -100.7866
24.8529 | -100.8448
24.8565 | -100.9026
24.8602 | -100.9591
24.8639 | -101.0147
24.8675 | -101.0684
24.8712 | -101.1215
24.8748 | -101.1723
24.8785 | -101.2211
24.8821 | -101.268
24.8858 | -101.3111
24.8894 | -101.353
24.8931 | -101.3936
24.8968 | -101.4311
24.9004 | -101.4646
24.9041 | -101.4978
24.9077 | -101.5283
24.9114 | -101.5558
24.915 | -101.5829
24.9187 | -101.6075
24.9223 | -101.6341
24.926 | -101.6643
24.9296 | -101.6944
24.9333 | -101.7191
24.937 | -101.7478
24.9406 | -101.784
24.9443 | -101.8165
24.9479 | -101.8357
24.9516 | -101.8431
24.9552 | -101.8492
24.9589 | -101.8518
24.9625 | -101.8555
24.9662 | -101.8612
24.9698 | -101.8686
24.9735 | -101.8742
24.9772 | -101.8802
24.9808 | -101.8871
24.9845 | -101.8923
24.9881 | -101.8977
24.9918 | -101.9053
24.9954 | -101.9141
24.9991 | -101.923
25.0027 | -101.9329
25.0064 | -101.9452
25.0101 | -101.9578
25.0137 | -101.9682
25.0174 | -101.979
25.021 | -101.9929
25.0247 | -102.0069
25.0283 | -102.021
25.032 | -102.0375
25.0356 | -102.0553
25.0393 | -102.0725
25.0429 | -102.0894
25.0466 | -102.1064
25.0503 | -102.1238
25.0539 | -102.141
25.0576 | -102.159
25.0612 | -102.1782
25.0649 | -102.1979
25.0685 | -102.2167
25.0722 | -102.2357
25.0758 | -102.2555
25.0795 | -102.2729
25.0831 | -102.2918
25.0868 | -102.3104
25.0905 | -102.3298
25.0941 | -102.3491
25.0978 | -102.3665
25.1014 | -102.3828
25.1051 | -102.397
25.1087 | -102.4116
25.1124 | -102.4255
25.116 | -102.4401
25.1197 | -102.4532
25.1234 | -102.4658
25.127 | -102.4779
25.1307 | -102.4868
25.1343 | -102.494
25.138 | -102.4999
25.1416 | -102.5077
25.1453 | -102.5128
25.1489 | -102.5152
25.1526 | -102.5169
25.1562 | -102.518
25.1599 | -102.5186
25.1636 | -102.5182
25.1672 | -102.5149
25.1709 | -102.5113
25.1745 | -102.5084
25.1782 | -102.5076
25.1818 | -102.5064
25.1855 | -102.5034
25.1891 | -102.4995
25.1928 | -102.4919
25.1964 | -102.486
25.2001 | -102.4798
25.2038 | -102.4702
25.2074 | -102.4605
25.2111 | -102.4513
25.2147 | -102.4435
25.2184 | -102.4319
25.222 | -102.4204
25.2257 | -102.4102
25.2293 | -102.4007
25.233 | -102.3902
25.2367 | -102.3787
25.2403 | -102.3689
25.244 | -102.3604
25.2476 | -102.3516
25.2513 | -102.3433
25.2549 | -102.3367
25.2586 | -102.329
25.2622 | -102.3227
25.2659 | -102.3164
25.2695 | -102.3092
25.2732 | -102.3005
25.2769 | -102.2909
25.2805 | -102.2823
25.2842 | -102.2733
25.2878 | -102.2655
25.2915 | -102.2586
25.2951 | -102.2525
25.2988 | -102.2467
25.3024 | -102.2403
25.3061 | -102.2386
25.3097 | -102.2326
25.3134 | -102.227
25.3171 | -102.2223
25.3207 | -102.2184
25.3244 | -102.2155
25.328 | -102.2125
25.3317 | -102.2087
25.3353 | -102.2057
25.339 | -102.2048
25.3426 | -102.2037
25.3463 | -102.2042
25.35 | -102.2039
25.3536 | -102.2037
25.3573 | -102.2031
25.3609 | -102.2054
25.3646 | -102.2078
25.3682 | -102.2089
25.3719 | -102.2109
25.3755 | -102.2141
25.3792 | -102.2185
25.3828 | -102.2225
25.3865 | -102.2269
25.3902 | -102.2332
25.3938 | -102.2397
25.3975 | -102.2451
25.4011 | -102.2513
25.4048 | -102.258
25.4084 | -102.263
25.4121 | -102.2701
25.4157 | -102.2782
25.4194 | -102.2851
25.423 | -102.2904
25.4267 | -102.2972
25.4304 | -102.3034
25.434 | -102.3095
25.4377 | -102.318
25.4413 | -102.3272
25.445 | -102.3357
25.4486 | -102.3429
25.4523 | -102.3505
25.4559 | -102.3575
25.4596 | -102.363
25.4633 | -102.3689
25.4669 | -102.3741
25.4706 | -102.3797
25.4742 | -102.3843
25.4779 | -102.3882
25.4815 | -102.391
25.4852 | -102.3947
25.4888 | -102.3995
25.4925 | -102.4053
25.4961 | -102.4117
25.4998 | -102.4149
25.5035 | -102.418
25.5071 | -102.4215
25.5108 | -102.4241
25.5144 | -102.4274
25.5181 | -102.4295
25.5217 | -102.431
25.5254 | -102.4347
25.529 | -102.4364
25.5327 | -102.4366
25.5363 | -102.4358
25.54 | -102.4365
25.5437 | -102.437
25.5473 | -102.4389
25.551 | -102.4422
25.5546 | -102.4421
25.5583 | -102.442
25.5619 | -102.4433
25.5656 | -102.4442
25.5692 | -102.4449
25.5729 | -102.4426
25.5766 | -102.4425
25.5802 | -102.4425
25.5839 | -102.4432
25.5875 | -102.4441
25.5912 | -102.4456
25.5948 | -102.4459
25.5985 | -102.4435
25.6021 | -102.4396
25.6058 | -102.4361
25.6094 | -102.4317
25.6131 | -102.429
25.6168 | -102.4265
25.6204 | -102.4234
25.6241 | -102.4205
25.6277 | -102.4181
25.6314 | -102.4178
25.635 | -102.4156
25.6387 | -102.4115
25.6423 | -102.4083
25.646 | -102.4044
25.6496 | -102.4008
25.6533 | -102.3951
25.657 | -102.3908
25.6606 | -102.3879
25.6643 | -102.3846
25.6679 | -102.3809
25.6716 | -102.3763
25.6752 | -102.3717
25.6789 | -102.3692
25.6825 | -102.3659
25.6862 | -102.3605
25.6899 | -102.3575
25.6935 | -102.3558
25.6972 | -102.3534
25.7008 | -102.3508
25.7045 | -102.3494
25.7081 | -102.3473
25.7118 | -102.3466
25.7154 | -102.3473
25.7191 | -102.3488
25.7227 | -102.3491
25.7264 | -102.3502
25.7301 | -102.3533
25.7337 | -102.3584
25.7374 | -102.3645
25.741 | -102.3708
25.7447 | -102.3776
25.7483 | -102.3857
25.752 | -102.3929
25.7556 | -102.4002
25.7593 | -102.4098
25.7629 | -102.4197
25.7666 | -102.4313
25.7703 | -102.444
25.7739 | -102.4539
25.7776 | -102.4648
25.7812 | -102.4781
25.7849 | -102.4935
25.7885 | -102.5089
25.7922 | -102.524
25.7958 | -102.5394
25.7995 | -102.5533
25.8031 | -102.5677
25.8068 | -102.5828
25.8105 | -102.598
25.8141 | -102.613
25.8178 | -102.6263
25.8214 | -102.639
25.8251 | -102.6523
25.8287 | -102.6657
25.8324 | -102.6787
25.836 | -102.6855
25.8397 | -102.699
25.8434 | -102.7125
25.847 | -102.7264
25.8507 | -102.737
25.8543 | -102.7466
25.858 | -102.7552
25.8616 | -102.7614
25.8653 | -102.7684
25.8689 | -102.7734
25.8726 | -102.7776
25.8762 | -102.7829
25.8799 | -102.7867
25.8836 | -102.7895
25.8872 | -102.7915
25.8909 | -102.7918
25.8945 | -102.7919
25.8982 | -102.7929
25.9018 | -102.7934
25.9055 | -102.7916
25.9091 | -102.7892
25.9128 | -102.785
25.9165 | -102.7805
25.9201 | -102.7768
25.9238 | -102.7706
25.9274 | -102.7631
25.9311 | -102.7558
25.9347 | -102.7496
25.9384 | -102.7435
25.942 | -102.7373
25.9457 | -102.729
25.9493 | -102.7173
25.953 | -102.7059
25.9567 | -102.6946
25.9603 | -102.6851
25.964 | -102.6776
25.9676 | -102.6699
25.9713 | -102.6619
25.9749 | -102.6551
25.9786 | -102.6476
25.9822 | -102.6399
25.9859 | -102.6331
25.9895 | -102.6269
25.9932 | -102.6208
25.9969 | -102.6138
26.0005 | -102.6056
26.0042 | -102.6009
26.0078 | -102.5954
26.0115 | -102.5903
26.0151 | -102.587
26.0188 | -102.5857
26.0224 | -102.5843
26.0261 | -102.5827
26.0297 | -102.5796
26.0334 | -102.577
26.0371 | -102.5778
26.0407 | -102.5792
26.0444 | -102.578
26.048 | -102.5782
26.0517 | -102.5811
26.0553 | -102.5822
26.059 | -102.5827
26.0626 | -102.5843
26.0663 | -102.5873
26.07 | -102.5888
26.0736 | -102.5909
26.0773 | -102.5937
26.0809 | -102.597
26.0846 | -102.6006
26.0882 | -102.6034
26.0919 | -102.6043
26.0955 | -102.6066
26.0992 | -102.6088
26.1028 | -102.6118
26.1065 | -102.6156
26.1102 | -102.6204
26.1138 | -102.6238
26.1175 | -102.627
26.1211 | -102.6316
26.1248 | -102.6371
26.1284 | -102.6423
26.1321 | -102.6461
26.1357 | -102.6512
26.1394 | -102.6576
26.143 | -102.6612
26.1467 | -102.6651
26.1504 | -102.6706
26.154 | -102.6755
26.1577 | -102.6802
26.1613 | -102.684
26.165 | -102.689
26.1686 | -102.6933
26.1723 | -102.6971
26.1759 | -102.7018
26.1796 | -102.7071
26.1833 | -102.7138
26.1869 | -102.7206
26.1906 | -102.726
26.1942 | -102.7307
26.1979 | -102.7362
26.2015 | -102.7422
26.2052 | -102.7468
26.2088 | -102.7503
26.2125 | -102.7544
26.2161 | -102.7574
26.2198 | -102.7605
26.2235 | -102.762
26.2271 | -102.7654
26.2308 | -102.7704
26.2344 | -102.7744
26.2381 | -102.7786
26.2417 | -102.7818
26.2454 | -102.7819
26.249 | -102.7828
26.2527 | -102.786
26.2564 | -102.788
26.26 | -102.7878
26.2637 | -102.7875
26.2673 | -102.7867
26.271 | -102.7846
26.2746 | -102.7826
26.2783 | -102.7808
26.2819 | -102.7805
26.2856 | -102.7804
26.2892 | -102.779
26.2929 | -102.7793
26.2966 | -102.7792
26.3002 | -102.7769
26.3039 | -102.7729
26.3075 | -102.7683
26.3112 | -102.7603
26.3148 | -102.752
26.3185 | -102.7465
26.3221 | -102.7401
26.3258 | -102.7328
26.3294 | -102.7231
26.3331 | -102.7121
26.3368 | -102.7005
26.3404 | -102.6897
26.3441 | -102.6779
26.3477 | -102.6648
26.3514 | -102.652
26.355 | -102.6391
26.3587 | -102.6263
26.3623 | -102.6063
26.366 | -102.5904
26.3697 | -102.5759
26.3733 | -102.5636
26.377 | -102.548
26.3806 | -102.5337
26.3843 | -102.5199
26.3879 | -102.5037
26.3916 | -102.4877
26.3952 | -102.4702
26.3989 | -102.4534
26.4025 | -102.4393
26.4062 | -102.426
26.4099 | -102.4123
26.4135 | -102.3978
26.4172 | -102.3833
26.4208 | -102.3696
26.4245 | -102.3566
26.4281 | -102.3427
26.4318 | -102.3301
26.4354 | -102.3164
26.4391 | -102.3004
26.4427 | -102.2873
26.4464 | -102.2752
26.4501 | -102.2652
26.4537 | -102.2558
26.4574 | -102.2465
26.461 | -102.2372
26.4647 | -102.2287
26.4683 | -102.2181
26.472 | -102.209
26.4756 | -102.1997
26.4793 | -102.1892
26.4829 | -102.1793
26.4866 | -102.1729
26.4903 | -102.1663
26.4939 | -102.1581
26.4976 | -102.1495
26.5012 | -102.1422
26.5049 | -102.1357
26.5085 | -102.1313
26.5122 | -102.1264
26.5158 | -102.1212
26.5195 | -102.1147
26.5232 | -102.1079
26.5268 | -102.1021
26.5305 | -102.0971
26.5341 | -102.0928
26.5378 | -102.0869
26.5414 | -102.0793
26.5451 | -102.0721
26.5487 | -102.0651
26.5524 | -102.0559
26.556 | -102.0482
26.5597 | -102.0414
26.5634 | -102.035
26.567 | -102.0276
26.5707 | -102.0193
26.5743 | -102.0103
26.578 | -102.0015
26.5816 | -101.9931
26.5853 | -101.9843
26.5889 | -101.9765
26.5926 | -101.9693
26.5963 | -101.9623
26.5999 | -101.9527
26.6036 | -101.9405
26.6072 | -101.9298
26.6109 | -101.9197
26.6145 | -101.9087
26.6182 | -101.8996
26.6218 | -101.8886
26.6255 | -101.8766
26.6291 | -101.8671
26.6328 | -101.8577
26.6365 | -101.8471
26.6401 | -101.8377
26.6438 | -101.8294
26.6474 | -101.82
26.6511 | -101.8099
26.6547 | -101.7983
26.6584 | -101.7867
26.662 | -101.7757
26.6657 | -101.7643
26.6693 | -101.7538
26.673 | -101.7445
26.6767 | -101.7356
26.6803 | -101.7281
26.684 | -101.7209
26.6876 | -101.7142
26.6913 | -101.7067
26.6949 | -101.6998
26.6986 | -101.6925
26.7022 | -101.684
26.7059 | -101.6775
26.7096 | -101.6718
26.7132 | -101.6659
26.7169 | -101.6593
26.7205 | -101.6526
26.7242 | -101.6475
26.7278 | -101.6432
26.7315 | -101.6372
26.7351 | -101.633
26.7388 | -101.631
26.7424 | -101.6274
26.7461 | -101.6236
26.7498 | -101.6205
26.7534 | -101.6179
26.7571 | -101.6143
26.7607 | -101.6109
26.7644 | -101.6096
26.768 | -101.6091
26.7717 | -101.6086
26.7753 | -101.6085
26.779 | -101.6078
26.7826 | -101.6081
26.7863 | -101.6083
26.79 | -101.6086
26.7936 | -101.611
26.7973 | -101.6134
26.8009 | -101.6155
26.8046 | -101.6194
26.8082 | -101.6225
26.8119 | -101.6235
26.8155 | -101.6282
26.8192 | -101.6334
26.8228 | -101.6363
26.8265 | -101.6378
26.8302 | -101.639
26.8338 | -101.6415
26.8375 | -101.646
26.8411 | -101.6522
26.8448 | -101.6586
26.8484 | -101.6637
26.8521 | -101.669
26.8557 | -101.6733
26.8594 | -101.6775
26.8631 | -101.6794
26.8667 | -101.6814
26.8704 | -101.6855
26.874 | -101.6884
26.8777 | -101.6891
26.8813 | -101.6896
26.885 | -101.6872
26.8886 | -101.6843
26.8923 | -101.6846
26.8959 | -101.6849
26.8996 | -101.6836
26.9033 | -101.6811
26.9069 | -101.6771
26.9106 | -101.671
26.9142 | -101.6623
26.9179 | -101.6504
26.9215 | -101.6414
26.9252 | -101.6345
26.9288 | -101.6257
26.9325 | -101.6176
26.9362 | -101.6076
26.9398 | -101.5965
26.9435 | -101.5852
26.9471 | -101.5745
26.9508 | -101.5631
26.9544 | -101.5516
26.9581 | -101.5389
26.9617 | -101.5256
26.9654 | -101.511
26.969 | -101.4942
26.9727 | -101.4775
26.9764 | -101.4623
26.98 | -101.448
26.9837 | -101.4345
26.9873 | -101.4186
26.991 | -101.4006
26.9946 | -101.385
26.9983 | -101.3689
27.0019 | -101.3524
27.0056 | -101.336
27.0092 | -101.3184
27.0129 | -101.3019
27.0166 | -101.2863
27.0202 | -101.2687
27.0239 | -101.2506
27.0275 | -101.2338
27.0312 | -101.2177
27.0348 | -101.2014
27.0385 | -101.1857
27.0421 | -101.1708
27.0458 | -101.1552
27.0495 | -101.1402
27.0531 | -101.1249
27.0568 | -101.1103
27.0604 | -101.0971
27.0641 | -101.0826
27.0677 | -101.0677
27.0714 | -101.0544
27.075 | -101.0425
27.0787 | -101.0328
27.0823 | -101.0247
27.086 | -101.014
27.0897 | -101.0033
27.0933 | -100.9927
27.097 | -100.983
27.1006 | -100.9716
27.1043 | -100.956
27.1079 | -100.9412
27.1116 | -100.9291
27.1152 | -100.9167
27.1189 | -100.9061
27.1225 | -100.8984
27.1262 | -100.8869
27.1299 | -100.872
27.1335 | -100.8582
27.1372 | -100.8465
27.1408 | -100.8347
27.1445 | -100.823
27.1481 | -100.8105
27.1518 | -100.8001
27.1554 | -100.789
27.1591 | -100.7779
27.1627 | -100.7658
27.1664 | -100.7529
27.1701 | -100.7398
27.1737 | -100.7253
27.1774 | -100.7097
27.181 | -100.6958
27.1847 | -100.6814
27.1883 | -100.6661
27.192 | -100.6516
27.1956 | -100.636
27.1993 | -100.6192
27.203 | -100.6038
27.2066 | -100.5884
27.2103 | -100.5735
27.2139 | -100.556
27.2176 | -100.5373
27.2212 | -100.5189
27.2249 | -100.5008
27.2285 | -100.4819
27.2322 | -100.4635
27.2358 | -100.4466
27.2395 | -100.4313
27.2432 | -100.4164
27.2468 | -100.4013
27.2505 | -100.3849
27.2541 | -100.3696
27.2578 | -100.3564
27.2614 | -100.343
27.2651 | -100.3291
27.2687 | -100.3137
27.2724 | -100.2971
27.2761 | -100.2824
27.2797 | -100.2697
27.2834 | -100.2573
27.287 | -100.2462
27.2907 | -100.2383
27.2943 | -100.2332
27.298 | -100.2303
27.3016 | -100.2291
27.3053 | -100.2255
27.3089 | -100.2212
27.3126 | -100.2192
27.3163 | -100.2179
27.3199 | -100.2171
27.3236 | -100.2161
27.3272 | -100.2148
27.3309 | -100.2135
27.3345 | -100.2138
27.3382 | -100.2149
27.3418 | -100.2181
27.3455 | -100.2221
27.3491 | -100.2268
27.3528 | -100.233
27.3565 | -100.2387
27.3601 | -100.2442
27.3638 | -100.2505
27.3674 | -100.2588
27.3711 | -100.2677
27.3747 | -100.276
27.3784 | -100.2832
27.382 | -100.291
27.3857 | -100.299
27.3894 | -100.3072
27.393 | -100.3161
27.3967 | -100.3231
27.4003 | -100.33
27.404 | -100.3375
27.4076 | -100.3441
27.4113 | -100.3488
27.4149 | -100.353
27.4186 | -100.3583
27.4222 | -100.3642
27.4259 | -100.3713
27.4296 | -100.3788
27.4332 | -100.386
27.4369 | -100.3919
27.4405 | -100.3973
27.4442 | -100.4026
27.4478 | -100.4081
27.4515 | -100.415
27.4551 | -100.4217
27.4588 | -100.4255
27.4624 | -100.4279
27.4661 | -100.43
27.4698 | -100.4319
27.4734 | -100.4357
27.4771 | -100.4399
27.4807 | -100.4444
27.4844 | -100.4486
27.488 | -100.4543
27.4917 | -100.4638
27.4953 | -100.4707
27.499 | -100.476
27.5026 | -100.481
27.5063 | -100.4838
27.51 | -100.4858
27.5136 | -100.4921
27.5173 | -100.5008
27.5209 | -100.5071
27.5246 | -100.5115
27.5282 | -100.5165
27.5319 | -100.5211
27.5355 | -100.5267
27.5392 | -100.5312
27.5429 | -100.5334
27.5465 | -100.5373
27.5502 | -100.5423
27.5538 | -100.5475
27.5575 | -100.5545
27.5611 | -100.5615
27.5648 | -100.5693
27.5684 | -100.5767
27.5721 | -100.5838
27.5757 | -100.5918
27.5794 | -100.601
27.5831 | -100.6082
27.5867 | -100.6153
27.5904 | -100.6231
27.594 | -100.6319
27.5977 | -100.641
27.6013 | -100.6518
27.605 | -100.6637
27.6086 | -100.6747
27.6123 | -100.6848
27.616 | -100.6931
27.6196 | -100.7008
27.6233 | -100.7087
27.6269 | -100.7179
27.6306 | -100.7282
27.6342 | -100.7385
27.6379 | -100.7496
27.6415 | -100.7616
27.6452 | -100.7724
27.6488 | -100.7835
27.6525 | -100.793
27.6562 | -100.8026
27.6598 | -100.815
27.6635 | -100.8263
27.6671 | -100.8377
27.6708 | -100.8478
27.6744 | -100.8559
27.6781 | -100.8625
27.6817 | -100.8687
27.6854 | -100.8742
27.689 | -100.8808
27.6927 | -100.8885
27.6964 | -100.8944
27.7 | -100.8996
27.7037 | -100.9029
27.7073 | -100.9037
27.711 | -100.9065
27.7146 | -100.9104
27.7183 | -100.914
27.7219 | -100.9161
27.7256 | -100.9171
27.7293 | -100.9152
27.7329 | -100.9134
27.7366 | -100.9129
27.7402 | -100.9104
27.7439 | -100.9085
27.7475 | -100.9062
27.7512 | -100.9007
27.7548 | -100.895
27.7585 | -100.8889
27.7621 | -100.8827
27.7658 | -100.8766
27.7695 | -100.8691
27.7731 | -100.8619
27.7768 | -100.8545
27.7804 | -100.8469
27.7841 | -100.8393
27.7877 | -100.832
27.7914 | -100.8232
27.795 | -100.813
27.7987 | -100.8031
27.8023 | -100.7911
27.806 | -100.7769
27.8097 | -100.767
27.8133 | -100.7571
27.817 | -100.7448
27.8206 | -100.7331
27.8243 | -100.7233
27.8279 | -100.7145
27.8316 | -100.7055
27.8352 | -100.6962
27.8389 | -100.6858
27.8425 | -100.6765
27.8462 | -100.668
27.8499 | -100.6605
27.8535 | -100.6536
27.8572 | -100.6444
27.8608 | -100.6356
27.8645 | -100.6279
27.8681 | -100.6199
27.8718 | -100.6111
27.8754 | -100.6039
27.8791 | -100.5973
27.8828 | -100.59
27.8864 | -100.5836
27.8901 | -100.5772
27.8937 | -100.571
27.8974 | -100.5652
27.901 | -100.5586
27.9047 | -100.5516
27.9083 | -100.5448
27.912 | -100.5396
27.9156 | -100.5362
27.9193 | -100.5323
27.923 | -100.5279
27.9266 | -100.5238
27.9303 | -100.5206
27.9339 | -100.5186
27.9376 | -100.514
27.9412 | -100.5095
27.9449 | -100.5071
27.9485 | -100.5041
27.9522 | -100.5007
27.9559 | -100.4974
27.9595 | -100.4921
27.9632 | -100.4847
27.9668 | -100.4741
27.9705 | -100.4656
27.9741 | -100.4604
27.9778 | -100.4573
27.9814 | -100.454
27.9851 | -100.4503
27.9887 | -100.4466
27.9924 | -100.4422
27.9961 | -100.4376
27.9997 | -100.4311
28.0034 | -100.4236