## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3d | 30 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 24
Hood | 02-91C/P | 15
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1001.2 | hPa |  | PRC PRO 01
Humidity | 31.6 | %/Hr
Internal temperature of calibration bench | 23.54 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1009 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 23.2 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 1.99 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.82 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.01 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3150 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Analog acquisition system | Analog Hood
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 28/02/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 09/03/2018
Technical Manager | Anthony HUE | Approval |  | 20/03/2018
CEO | Robert GLEMOT | Approval |  | 20/03/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4302 | 57.5986
0.012209 | -35.973 | 38.5279
0.018313 | -34.9288 | 28.5478
0.024417 | -34.4753 | 22.1485
0.030521 | -34.3612 | 18.3339
0.036626 | -34.2177 | 15.2416
0.04273 | -34.132 | 12.9381
0.048834 | -34.0794 | 11.3451
0.054938 | -34.0117 | 10.1387
0.061043 | -33.9405 | 9.0215
0.067147 | -33.9295 | 8.1556
0.073251 | -33.9291 | 7.4027
0.079355 | -33.9444 | 6.8353
0.08546 | -33.9598 | 6.3657
0.091564 | -33.9557 | 5.9801
0.097668 | -33.9241 | 5.711
0.10377 | -33.8803 | 5.3681
0.10988 | -33.8517 | 5.0103
0.11598 | -33.8429 | 4.656
0.12209 | -33.8578 | 4.375
0.12819 | -33.8872 | 4.1369
0.13429 | -33.9156 | 3.8946
0.1404 | -33.9308 | 3.6282
0.1465 | -33.9448 | 3.3758
0.15261 | -33.9512 | 3.1425
0.15871 | -33.9482 | 2.9497
0.16482 | -33.9353 | 2.7918
0.17092 | -33.913 | 2.7222
0.17702 | -33.8957 | 2.6208
0.18313 | -33.8799 | 2.5332
0.18923 | -33.8684 | 2.4255
0.19534 | -33.862 | 2.3209
0.20144 | -33.861 | 2.2081
0.20754 | -33.8637 | 2.1318
0.21365 | -33.8659 | 2.0543
0.21975 | -33.8698 | 1.9861
0.22586 | -33.8769 | 1.9177
0.23196 | -33.8845 | 1.8425
0.23807 | -33.8921 | 1.7631
0.24417 | -33.8938 | 1.6711
0.25027 | -33.8966 | 1.5872
0.25638 | -33.8963 | 1.4967
0.26248 | -33.8944 | 1.4113
0.26859 | -33.8921 | 1.3294
0.27469 | -33.8913 | 1.2448
0.2808 | -33.8965 | 1.1801
0.2869 | -33.9004 | 1.1115
0.293 | -33.9053 | 1.0531
0.29911 | -33.9109 | 1.0115
0.30521 | -33.9163 | 0.97162
0.31132 | -33.92 | 0.95255
0.31742 | -33.9239 | 0.92589
0.32353 | -33.927 | 0.90317
0.32963 | -33.9296 | 0.8834
0.33573 | -33.9314 | 0.86573
0.34184 | -33.9337 | 0.84531
0.34794 | -33.9365 | 0.80878
0.35405 | -33.9394 | 0.77662
0.36015 | -33.9425 | 0.7415
0.36626 | -33.9446 | 0.70331
0.37236 | -33.9467 | 0.65375
0.37846 | -33.9486 | 0.59764
0.38457 | -33.9477 | 0.53995
0.39067 | -33.9482 | 0.49016
0.39678 | -33.9466 | 0.43492
0.40288 | -33.9444 | 0.38257
0.40899 | -33.9409 | 0.33133
0.41509 | -33.9368 | 0.27985
0.42119 | -33.9322 | 0.23293
0.4273 | -33.9287 | 0.18448
0.4334 | -33.9255 | 0.13537
0.43951 | -33.9227 | 0.08711
0.44561 | -33.9212 | 0.041753
0.45172 | -33.9225 | 0.0037005
0.45782 | -33.9222 | -0.038274
0.46392 | -33.9232 | -0.077076
0.47003 | -33.9247 | -0.11271
0.47613 | -33.9265 | -0.14554
0.48224 | -33.9285 | -0.17691
0.48834 | -33.9308 | -0.1999
0.49445 | -33.9327 | -0.226
0.50055 | -33.935 | -0.2499
0.50665 | -33.9366 | -0.26995
0.51276 | -33.9378 | -0.29134
0.51886 | -33.9389 | -0.31043
0.52497 | -33.9381 | -0.31965
0.53107 | -33.9386 | -0.33228
0.53717 | -33.939 | -0.34577
0.54328 | -33.9389 | -0.35381
0.54938 | -33.9389 | -0.36457
0.55549 | -33.9385 | -0.37611
0.56159 | -33.9385 | -0.38976
0.5677 | -33.9381 | -0.40282
0.5738 | -33.9377 | -0.4184
0.5799 | -33.9378 | -0.43296
0.58601 | -33.9382 | -0.45164
0.59211 | -33.9392 | -0.47595
0.59822 | -33.9398 | -0.49484
0.60432 | -33.9407 | -0.51809
0.61043 | -33.9416 | -0.54117
0.61653 | -33.9427 | -0.56446
0.62263 | -33.9437 | -0.5891
0.62874 | -33.9449 | -0.61104
0.63484 | -33.946 | -0.63385
0.64095 | -33.9471 | -0.65649
0.64705 | -33.9482 | -0.67816
0.65316 | -33.9491 | -0.69925
0.65926 | -33.9501 | -0.723
0.66536 | -33.9496 | -0.74269
0.67147 | -33.9498 | -0.76073
0.67757 | -33.9503 | -0.78017
0.68368 | -33.9502 | -0.80137
0.68978 | -33.9503 | -0.82077
0.69589 | -33.9499 | -0.84034
0.70199 | -33.9492 | -0.86303
0.70809 | -33.9487 | -0.88433
0.7142 | -33.9483 | -0.90476
0.7203 | -33.9475 | -0.92545
0.72641 | -33.9468 | -0.94726
0.73251 | -33.9466 | -0.96901
0.73862 | -33.9466 | -0.994
0.74472 | -33.9468 | -1.0145
0.75082 | -33.9464 | -1.0364
0.75693 | -33.9465 | -1.0606
0.76303 | -33.9465 | -1.084
0.76914 | -33.9469 | -1.1114
0.77524 | -33.9471 | -1.137
0.78135 | -33.9475 | -1.1611
0.78745 | -33.948 | -1.1865
0.79355 | -33.9485 | -1.2117
0.79966 | -33.9494 | -1.238
0.80576 | -33.9504 | -1.2684
0.81187 | -33.9507 | -1.298
0.81797 | -33.9512 | -1.3287
0.82408 | -33.9518 | -1.3562
0.83018 | -33.9523 | -1.3872
0.83628 | -33.9531 | -1.4177
0.84239 | -33.953 | -1.4458
0.84849 | -33.953 | -1.474
0.8546 | -33.9534 | -1.5018
0.8607 | -33.9534 | -1.531
0.86681 | -33.9535 | -1.5594
0.87291 | -33.9531 | -1.5854
0.87901 | -33.9527 | -1.6083
0.88512 | -33.9525 | -1.6318
0.89122 | -33.9524 | -1.6556
0.89733 | -33.9522 | -1.6757
0.90343 | -33.9519 | -1.6976
0.90953 | -33.9521 | -1.7137
0.91564 | -33.9516 | -1.7308
0.92174 | -33.9513 | -1.7463
0.92785 | -33.9509 | -1.7612
0.93395 | -33.9509 | -1.7744
0.94006 | -33.9509 | -1.7894
0.94616 | -33.9508 | -1.7962
0.95226 | -33.9509 | -1.8093
0.95837 | -33.951 | -1.8225
0.96447 | -33.9514 | -1.8344
0.97058 | -33.9514 | -1.846
0.97668 | -33.9512 | -1.8558
0.98279 | -33.9513 | -1.8617
0.98889 | -33.9515 | -1.8687
0.99499 | -33.9516 | -1.879
1.0011 | -33.9519 | -1.8861
1.0072 | -33.9518 | -1.8921
1.0133 | -33.9518 | -1.8982
1.0194 | -33.9518 | -1.9057
1.0255 | -33.9515 | -1.9145
1.0316 | -33.951 | -1.9208
1.0377 | -33.9506 | -1.9288
1.0438 | -33.9499 | -1.9354
1.0499 | -33.9486 | -1.9438
1.056 | -33.9476 | -1.9507
1.0621 | -33.9466 | -1.9585
1.0682 | -33.9462 | -1.9656
1.0743 | -33.9452 | -1.9752
1.0805 | -33.9446 | -1.984
1.0866 | -33.9432 | -1.9963
1.0927 | -33.942 | -2.0056
1.0988 | -33.9413 | -2.0178
1.1049 | -33.94 | -2.0272
1.111 | -33.9386 | -2.0399
1.1171 | -33.9374 | -2.051
1.1232 | -33.9361 | -2.0664
1.1293 | -33.9351 | -2.0798
1.1354 | -33.9343 | -2.0936
1.1415 | -33.9332 | -2.1089
1.1476 | -33.932 | -2.1241
1.1537 | -33.9308 | -2.1387
1.1598 | -33.9307 | -2.1579
1.1659 | -33.9302 | -2.1714
1.172 | -33.9296 | -2.186
1.1781 | -33.9289 | -2.2013
1.1842 | -33.9284 | -2.218
1.1903 | -33.9291 | -2.2344
1.1964 | -33.9291 | -2.2501
1.2025 | -33.929 | -2.2665
1.2086 | -33.9295 | -2.2833
1.2147 | -33.9297 | -2.2995
1.2209 | -33.9304 | -2.3143
1.227 | -33.9321 | -2.3305
1.2331 | -33.933 | -2.346
1.2392 | -33.9339 | -2.3629
1.2453 | -33.9351 | -2.3764
1.2514 | -33.9363 | -2.39
1.2575 | -33.9377 | -2.4053
1.2636 | -33.9399 | -2.4182
1.2697 | -33.9415 | -2.4314
1.2758 | -33.9429 | -2.4468
1.2819 | -33.9446 | -2.4616
1.288 | -33.9461 | -2.4766
1.2941 | -33.948 | -2.4888
1.3002 | -33.9499 | -2.4989
1.3063 | -33.9514 | -2.5105
1.3124 | -33.9531 | -2.5221
1.3185 | -33.9549 | -2.5343
1.3246 | -33.9566 | -2.5466
1.3307 | -33.9586 | -2.5578
1.3368 | -33.9603 | -2.5671
1.3429 | -33.9618 | -2.5776
1.349 | -33.9631 | -2.5878
1.3551 | -33.9647 | -2.6002
1.3613 | -33.9661 | -2.6099
1.3674 | -33.9673 | -2.62
1.3735 | -33.9685 | -2.6298
1.3796 | -33.9697 | -2.639
1.3857 | -33.9706 | -2.6478
1.3918 | -33.9716 | -2.6592
1.3979 | -33.9728 | -2.6694
1.404 | -33.9734 | -2.6809
1.4101 | -33.9743 | -2.6915
1.4162 | -33.9751 | -2.7018
1.4223 | -33.9758 | -2.7136
1.4284 | -33.9765 | -2.7236
1.4345 | -33.9771 | -2.7354
1.4406 | -33.9776 | -2.748
1.4467 | -33.9782 | -2.759
1.4528 | -33.9788 | -2.7698
1.4589 | -33.9796 | -2.7818
1.465 | -33.9798 | -2.7913
1.4711 | -33.9807 | -2.8019
1.4772 | -33.9812 | -2.8153
1.4833 | -33.9818 | -2.8289
1.4894 | -33.9824 | -2.8417
1.4955 | -33.9831 | -2.8546
1.5016 | -33.9837 | -2.8679
1.5078 | -33.9841 | -2.8836
1.5139 | -33.9848 | -2.8975
1.52 | -33.9853 | -2.9119
1.5261 | -33.9857 | -2.9249
1.5322 | -33.9859 | -2.9387
1.5383 | -33.9862 | -2.9522
1.5444 | -33.9868 | -2.9667
1.5505 | -33.9872 | -2.9811
1.5566 | -33.9877 | -2.9951
1.5627 | -33.9882 | -3.0096
1.5688 | -33.989 | -3.024
1.5749 | -33.9892 | -3.0375
1.581 | -33.9896 | -3.055
1.5871 | -33.9901 | -3.0715
1.5932 | -33.9905 | -3.0856
1.5993 | -33.9908 | -3.1032
1.6054 | -33.9912 | -3.1191
1.6115 | -33.9915 | -3.1369
1.6176 | -33.9919 | -3.154
1.6237 | -33.9921 | -3.1695
1.6298 | -33.9924 | -3.1858
1.6359 | -33.9928 | -3.2014
1.642 | -33.9932 | -3.2169
1.6482 | -33.9936 | -3.2347
1.6543 | -33.9941 | -3.2523
1.6604 | -33.9945 | -3.2679
1.6665 | -33.9946 | -3.2838
1.6726 | -33.9949 | -3.3009
1.6787 | -33.9954 | -3.3187
1.6848 | -33.9952 | -3.336
1.6909 | -33.9956 | -3.3531
1.697 | -33.9958 | -3.3701
1.7031 | -33.9959 | -3.3868
1.7092 | -33.9959 | -3.4026
1.7153 | -33.9962 | -3.4197
1.7214 | -33.9958 | -3.4361
1.7275 | -33.9957 | -3.4528
1.7336 | -33.9955 | -3.4699
1.7397 | -33.9956 | -3.4861
1.7458 | -33.9956 | -3.5036
1.7519 | -33.9955 | -3.5206
1.758 | -33.9954 | -3.5366
1.7641 | -33.9955 | -3.5521
1.7702 | -33.9954 | -3.5688
1.7763 | -33.9954 | -3.5836
1.7824 | -33.9953 | -3.6007
1.7885 | -33.9951 | -3.6159
1.7947 | -33.9952 | -3.6318
1.8008 | -33.9952 | -3.6459
1.8069 | -33.9951 | -3.6607
1.813 | -33.9952 | -3.6759
1.8191 | -33.9949 | -3.6913
1.8252 | -33.9947 | -3.7064
1.8313 | -33.9948 | -3.7219
1.8374 | -33.9948 | -3.7366
1.8435 | -33.9949 | -3.7501
1.8496 | -33.9947 | -3.7656
1.8557 | -33.9946 | -3.7815
1.8618 | -33.9946 | -3.7947
1.8679 | -33.9946 | -3.8089
1.874 | -33.9947 | -3.8233
1.8801 | -33.9946 | -3.8375
1.8862 | -33.9947 | -3.8516
1.8923 | -33.9947 | -3.8651
1.8984 | -33.9944 | -3.8773
1.9045 | -33.9944 | -3.8898
1.9106 | -33.9942 | -3.903
1.9167 | -33.9939 | -3.9157
1.9228 | -33.9939 | -3.928
1.9289 | -33.9937 | -3.9385
1.9351 | -33.9938 | -3.9492
1.9412 | -33.9938 | -3.961
1.9473 | -33.9939 | -3.9744
1.9534 | -33.9939 | -3.987
1.9595 | -33.9937 | -3.9999
1.9656 | -33.9933 | -4.012
1.9717 | -33.9932 | -4.0241
1.9778 | -33.9932 | -4.037
1.9839 | -33.9932 | -4.0498
1.99 | -33.9933 | -4.0612
1.9961 | -33.9931 | -4.0732
2.0022 | -33.9931 | -4.0847
2.0083 | -33.9928 | -4.0962
2.0144 | -33.9928 | -4.1076
2.0205 | -33.9929 | -4.1208
2.0266 | -33.993 | -4.1337
2.0327 | -33.9931 | -4.1463
2.0388 | -33.9934 | -4.1613
2.0449 | -33.9935 | -4.1731
2.051 | -33.9935 | -4.1846
2.0571 | -33.9935 | -4.1974
2.0632 | -33.9934 | -4.2094
2.0693 | -33.9934 | -4.222
2.0754 | -33.9937 | -4.2351
2.0816 | -33.9938 | -4.2484
2.0877 | -33.9938 | -4.2609
2.0938 | -33.9941 | -4.2746
2.0999 | -33.9941 | -4.2876
2.106 | -33.9943 | -4.3008
2.1121 | -33.9943 | -4.3143
2.1182 | -33.9945 | -4.3259
2.1243 | -33.9947 | -4.3379
2.1304 | -33.9948 | -4.3496
2.1365 | -33.9952 | -4.3625
2.1426 | -33.9953 | -4.377
2.1487 | -33.9958 | -4.39
2.1548 | -33.9962 | -4.4042
2.1609 | -33.9968 | -4.4175
2.167 | -33.9972 | -4.4313
2.1731 | -33.9976 | -4.4445
2.1792 | -33.9982 | -4.457
2.1853 | -33.9987 | -4.4711
2.1914 | -33.9992 | -4.4846
2.1975 | -33.9998 | -4.4979
2.2036 | -34.0002 | -4.5123
2.2097 | -34.0006 | -4.5256
2.2158 | -34.0015 | -4.5373
2.222 | -34.0021 | -4.5488
2.2281 | -34.0024 | -4.561
2.2342 | -34.0029 | -4.5753
2.2403 | -34.0033 | -4.5868
2.2464 | -34.0041 | -4.5993
2.2525 | -34.0045 | -4.6114
2.2586 | -34.0054 | -4.6228
2.2647 | -34.0062 | -4.6336
2.2708 | -34.0067 | -4.6466
2.2769 | -34.0076 | -4.6594
2.283 | -34.008 | -4.6702
2.2891 | -34.0086 | -4.682
2.2952 | -34.0091 | -4.6932
2.3013 | -34.0098 | -4.7048
2.3074 | -34.0104 | -4.716
2.3135 | -34.0109 | -4.7286
2.3196 | -34.0115 | -4.7387
2.3257 | -34.0119 | -4.7507
2.3318 | -34.0122 | -4.7612
2.3379 | -34.0127 | -4.7739
2.344 | -34.0131 | -4.7861
2.3501 | -34.0135 | -4.7983
2.3562 | -34.0138 | -4.8082
2.3623 | -34.0141 | -4.8184
2.3685 | -34.0144 | -4.829
2.3746 | -34.0147 | -4.8409
2.3807 | -34.0152 | -4.8514
2.3868 | -34.0154 | -4.8619
2.3929 | -34.0158 | -4.8727
2.399 | -34.016 | -4.8824
2.4051 | -34.0165 | -4.8934
2.4112 | -34.0169 | -4.9039
2.4173 | -34.0172 | -4.9144
2.4234 | -34.0172 | -4.9245
2.4295 | -34.0175 | -4.9346
2.4356 | -34.0177 | -4.9455
2.4417 | -34.0177 | -4.957
2.4478 | -34.0179 | -4.968
2.4539 | -34.0182 | -4.9775
2.46 | -34.0179 | -4.9887
2.4661 | -34.018 | -4.9995
2.4722 | -34.0183 | -5.0098
2.4783 | -34.0184 | -5.0198
2.4844 | -34.0185 | -5.0299
2.4905 | -34.0184 | -5.0407
2.4966 | -34.0182 | -5.0511
2.5027 | -34.0182 | -5.0608
2.5089 | -34.0182 | -5.0709
2.515 | -34.018 | -5.0817
2.5211 | -34.018 | -5.0916
2.5272 | -34.0178 | -5.1022
2.5333 | -34.0177 | -5.1132
2.5394 | -34.0177 | -5.1238
2.5455 | -34.0175 | -5.1347
2.5516 | -34.0175 | -5.145
2.5577 | -34.0175 | -5.1564
2.5638 | -34.0173 | -5.1682
2.5699 | -34.0172 | -5.1799
2.576 | -34.017 | -5.1917
2.5821 | -34.0169 | -5.2031
2.5882 | -34.0167 | -5.2141
2.5943 | -34.0167 | -5.2254
2.6004 | -34.0166 | -5.2368
2.6065 | -34.0164 | -5.2469
2.6126 | -34.0164 | -5.2575
2.6187 | -34.0163 | -5.2679
2.6248 | -34.0162 | -5.2785
2.6309 | -34.0161 | -5.2911
2.637 | -34.0162 | -5.3036
2.6431 | -34.0162 | -5.3146
2.6492 | -34.0161 | -5.3261
2.6554 | -34.0162 | -5.3378
2.6615 | -34.016 | -5.3491
2.6676 | -34.016 | -5.361
2.6737 | -34.0162 | -5.3731
2.6798 | -34.0162 | -5.3845
2.6859 | -34.0163 | -5.3964
2.692 | -34.0162 | -5.4076
2.6981 | -34.0162 | -5.4202
2.7042 | -34.0163 | -5.4335
2.7103 | -34.0164 | -5.445
2.7164 | -34.0166 | -5.4571
2.7225 | -34.0167 | -5.4697
2.7286 | -34.0168 | -5.4823
2.7347 | -34.0169 | -5.4937
2.7408 | -34.0172 | -5.5045
2.7469 | -34.0172 | -5.5172
2.753 | -34.0175 | -5.5292
2.7591 | -34.0177 | -5.5424
2.7652 | -34.0177 | -5.5553
2.7713 | -34.0178 | -5.5666
2.7774 | -34.0182 | -5.5793
2.7835 | -34.0184 | -5.5914
2.7896 | -34.0185 | -5.6022
2.7958 | -34.0189 | -5.614
2.8019 | -34.0195 | -5.6253
2.808 | -34.0198 | -5.6384
2.8141 | -34.0202 | -5.6525
2.8202 | -34.0206 | -5.6659
2.8263 | -34.021 | -5.6782
2.8324 | -34.0213 | -5.6919
2.8385 | -34.0217 | -5.7052
2.8446 | -34.0223 | -5.7182
2.8507 | -34.0227 | -5.7316
2.8568 | -34.0231 | -5.7441
2.8629 | -34.0236 | -5.7567
2.869 | -34.024 | -5.7694
2.8751 | -34.0244 | -5.7826
2.8812 | -34.0251 | -5.797
2.8873 | -34.0254 | -5.8097
2.8934 | -34.0258 | -5.8229
2.8995 | -34.0262 | -5.8361
2.9056 | -34.0265 | -5.8482
2.9117 | -34.027 | -5.8611
2.9178 | -34.0277 | -5.875
2.9239 | -34.0281 | -5.8895
2.93 | -34.0286 | -5.9041
2.9361 | -34.0291 | -5.9168
2.9423 | -34.0297 | -5.9301
2.9484 | -34.0302 | -5.9447
2.9545 | -34.0309 | -5.959
2.9606 | -34.0314 | -5.9735
2.9667 | -34.032 | -5.9866
2.9728 | -34.0324 | -6.0001
2.9789 | -34.0329 | -6.014
2.985 | -34.0333 | -6.0289
2.9911 | -34.0338 | -6.0426
2.9972 | -34.0343 | -6.0565
3.0033 | -34.0349 | -6.0707
3.0094 | -34.0355 | -6.0849
3.0155 | -34.0361 | -6.0987
3.0216 | -34.0366 | -6.1132
3.0277 | -34.037 | -6.128
3.0338 | -34.0376 | -6.1427
3.0399 | -34.0382 | -6.1575
3.046 | -34.0386 | -6.1717
3.0521 | -34.0392 | -6.1854
3.0582 | -34.0395 | -6.2003
3.0643 | -34.04 | -6.2145
3.0704 | -34.0403 | -6.229
3.0765 | -34.0407 | -6.2424
3.0827 | -34.0411 | -6.2566
3.0888 | -34.0414 | -6.2707
3.0949 | -34.0418 | -6.2848
3.101 | -34.0423 | -6.2999
3.1071 | -34.0426 | -6.3151
3.1132 | -34.0431 | -6.3302
3.1193 | -34.0435 | -6.3443
3.1254 | -34.0439 | -6.3577
3.1315 | -34.0444 | -6.3731
3.1376 | -34.0446 | -6.3871
3.1437 | -34.045 | -6.4017
3.1498 | -34.0454 | -6.4169
3.1559 | -34.0457 | -6.4313
3.162 | -34.0458 | -6.4463
3.1681 | -34.046 | -6.4616
3.1742 | -34.0464 | -6.4764
3.1803 | -34.0468 | -6.4907
3.1864 | -34.0471 | -6.5041
3.1925 | -34.0475 | -6.5188
3.1986 | -34.0476 | -6.5328
3.2047 | -34.0478 | -6.5478
3.2108 | -34.0481 | -6.5608
3.2169 | -34.0485 | -6.5755
3.223 | -34.0485 | -6.59
3.2292 | -34.0486 | -6.605
3.2353 | -34.0485 | -6.616
3.2414 | -34.0487 | -6.6301
3.2475 | -34.0488 | -6.6424
3.2536 | -34.0491 | -6.6572
3.2597 | -34.0493 | -6.6696
3.2658 | -34.0495 | -6.6817
3.2719 | -34.0497 | -6.6941
3.278 | -34.0498 | -6.7074
3.2841 | -34.05 | -6.7206
3.2902 | -34.0502 | -6.7342
3.2963 | -34.0503 | -6.7472
3.3024 | -34.0503 | -6.7587
3.3085 | -34.0506 | -6.7724
3.3146 | -34.0505 | -6.7846
3.3207 | -34.0508 | -6.798
3.3268 | -34.0509 | -6.8114
3.3329 | -34.0511 | -6.8226
3.339 | -34.0512 | -6.8337
3.3451 | -34.0513 | -6.8463
3.3512 | -34.0515 | -6.8572
3.3573 | -34.0517 | -6.8698
3.3634 | -34.0518 | -6.8826
3.3696 | -34.0518 | -6.8957
3.3757 | -34.0519 | -6.907
3.3818 | -34.0522 | -6.9199
3.3879 | -34.0522 | -6.9316
3.394 | -34.0522 | -6.9435
3.4001 | -34.0523 | -6.9556
3.4062 | -34.0524 | -6.9672
3.4123 | -34.0525 | -6.9783
3.4184 | -34.0526 | -6.9907
3.4245 | -34.0527 | -7.0025
3.4306 | -34.0527 | -7.0137
3.4367 | -34.0529 | -7.0248
3.4428 | -34.0532 | -7.0355
3.4489 | -34.0532 | -7.0478
3.455 | -34.0532 | -7.0596
3.4611 | -34.0533 | -7.072
3.4672 | -34.0533 | -7.0846
3.4733 | -34.0535 | -7.0961
3.4794 | -34.0535 | -7.1079
3.4855 | -34.0538 | -7.119
3.4916 | -34.0539 | -7.1301
3.4977 | -34.054 | -7.1423
3.5038 | -34.0542 | -7.1536
3.5099 | -34.0543 | -7.1647
3.5161 | -34.0544 | -7.1756
3.5222 | -34.0546 | -7.187
3.5283 | -34.0546 | -7.1985
3.5344 | -34.0547 | -7.2106
3.5405 | -34.0549 | -7.2219
3.5466 | -34.055 | -7.233
3.5527 | -34.0552 | -7.245
3.5588 | -34.0553 | -7.2562
3.5649 | -34.0554 | -7.2677
3.571 | -34.0555 | -7.2793
3.5771 | -34.0557 | -7.2914
3.5832 | -34.0559 | -7.3024
3.5893 | -34.0561 | -7.3136
3.5954 | -34.0563 | -7.3255
3.6015 | -34.0565 | -7.3372
3.6076 | -34.0567 | -7.3482
3.6137 | -34.0569 | -7.3595
3.6198 | -34.0572 | -7.3699
3.6259 | -34.0573 | -7.382
3.632 | -34.0576 | -7.3939
3.6381 | -34.0579 | -7.4052
3.6442 | -34.0581 | -7.417
3.6503 | -34.0584 | -7.4275
3.6565 | -34.0587 | -7.4404
3.6626 | -34.0588 | -7.4525
3.6687 | -34.0592 | -7.4644
3.6748 | -34.0595 | -7.4756
3.6809 | -34.0599 | -7.4875
3.687 | -34.0601 | -7.4999
3.6931 | -34.0605 | -7.5112
3.6992 | -34.0608 | -7.5232
3.7053 | -34.0611 | -7.5346
3.7114 | -34.0613 | -7.5472
3.7175 | -34.0617 | -7.5584
3.7236 | -34.0619 | -7.5687
3.7297 | -34.0621 | -7.5807
3.7358 | -34.0624 | -7.5936
3.7419 | -34.0626 | -7.6064
3.748 | -34.0628 | -7.6184
3.7541 | -34.063 | -7.6307
3.7602 | -34.0634 | -7.6428
3.7663 | -34.0637 | -7.6543
3.7724 | -34.064 | -7.6666
3.7785 | -34.0644 | -7.6783
3.7846 | -34.0646 | -7.6902
3.7907 | -34.0651 | -7.7025
3.7969 | -34.0653 | -7.7152
3.803 | -34.0657 | -7.7266
3.8091 | -34.0661 | -7.7382
3.8152 | -34.0664 | -7.7507
3.8213 | -34.0666 | -7.7635
3.8274 | -34.067 | -7.7753
3.8335 | -34.0675 | -7.7881
3.8396 | -34.0679 | -7.8008
3.8457 | -34.0683 | -7.8124
3.8518 | -34.0688 | -7.8245
3.8579 | -34.0692 | -7.8364
3.864 | -34.0696 | -7.8484
3.8701 | -34.0701 | -7.8606
3.8762 | -34.0704 | -7.873
3.8823 | -34.0707 | -7.8849
3.8884 | -34.0712 | -7.8966
3.8945 | -34.0717 | -7.9092
3.9006 | -34.0721 | -7.9219
3.9067 | -34.0725 | -7.9349
3.9128 | -34.0729 | -7.9474
3.9189 | -34.0733 | -7.9594
3.925 | -34.0736 | -7.9715
3.9311 | -34.0741 | -7.9829
3.9372 | -34.0745 | -7.9956
3.9434 | -34.0749 | -8.0074
3.9495 | -34.0753 | -8.0201
3.9556 | -34.0758 | -8.0326
3.9617 | -34.0761 | -8.0448
3.9678 | -34.0766 | -8.0567
3.9739 | -34.077 | -8.0692
3.98 | -34.0774 | -8.0822
3.9861 | -34.0777 | -8.0939
3.9922 | -34.078 | -8.1068
3.9983 | -34.0784 | -8.1182
4.0044 | -34.0788 | -8.1292
4.0105 | -34.0792 | -8.1414
4.0166 | -34.0797 | -8.1537
4.0227 | -34.0802 | -8.1653
4.0288 | -34.0804 | -8.1779
4.0349 | -34.0808 | -8.1892
4.041 | -34.0812 | -8.2014
4.0471 | -34.0815 | -8.2129
4.0532 | -34.0819 | -8.2253
4.0593 | -34.0823 | -8.2381
4.0654 | -34.0826 | -8.2502
4.0715 | -34.083 | -8.2628
4.0776 | -34.0834 | -8.2744
4.0838 | -34.0838 | -8.287
4.0899 | -34.0842 | -8.2995
4.096 | -34.0845 | -8.3112
4.1021 | -34.0849 | -8.3233
4.1082 | -34.0853 | -8.3348
4.1143 | -34.0857 | -8.3462
4.1204 | -34.0861 | -8.3582
4.1265 | -34.0865 | -8.3699
4.1326 | -34.087 | -8.3815
4.1387 | -34.0874 | -8.3939
4.1448 | -34.0878 | -8.4061
4.1509 | -34.0881 | -8.4181
4.157 | -34.0884 | -8.4306
4.1631 | -34.0887 | -8.443
4.1692 | -34.0891 | -8.455
4.1753 | -34.0895 | -8.4671
4.1814 | -34.0898 | -8.4783
4.1875 | -34.0902 | -8.4909
4.1936 | -34.0905 | -8.5035
4.1997 | -34.0909 | -8.5149
4.2058 | -34.0912 | -8.5277
4.2119 | -34.0915 | -8.5388
4.218 | -34.0918 | -8.5502
4.2241 | -34.0922 | -8.5612
4.2303 | -34.0927 | -8.5737
4.2364 | -34.093 | -8.5863
4.2425 | -34.0934 | -8.5981
4.2486 | -34.0937 | -8.6105
4.2547 | -34.0938 | -8.6211
4.2608 | -34.0942 | -8.6324
4.2669 | -34.0945 | -8.6441
4.273 | -34.0948 | -8.6562
4.2791 | -34.0951 | -8.6682
4.2852 | -34.0954 | -8.6798
4.2913 | -34.0956 | -8.6904
4.2974 | -34.0958 | -8.7026
4.3035 | -34.0961 | -8.7147
4.3096 | -34.0963 | -8.7264
4.3157 | -34.0966 | -8.7377
4.3218 | -34.0969 | -8.749
4.3279 | -34.0972 | -8.7607
4.334 | -34.0975 | -8.7719
4.3401 | -34.0977 | -8.7832
4.3462 | -34.0981 | -8.7952
4.3523 | -34.0985 | -8.8063
4.3584 | -34.0986 | -8.8163
4.3645 | -34.0989 | -8.8273
4.3707 | -34.0992 | -8.838
4.3768 | -34.0994 | -8.8497
4.3829 | -34.0996 | -8.8613
4.389 | -34.1 | -8.8724
4.3951 | -34.1002 | -8.8833
4.4012 | -34.1005 | -8.8945
4.4073 | -34.1008 | -8.9059
4.4134 | -34.101 | -8.9174
4.4195 | -34.1013 | -8.9284
4.4256 | -34.1017 | -8.9393
4.4317 | -34.102 | -8.9502
4.4378 | -34.1022 | -8.9611
4.4439 | -34.1024 | -8.972
4.45 | -34.1026 | -8.9833
4.4561 | -34.1028 | -8.9945
4.4622 | -34.103 | -9.0054
4.4683 | -34.1032 | -9.0161
4.4744 | -34.1034 | -9.0276
4.4805 | -34.1038 | -9.0381
4.4866 | -34.104 | -9.0487
4.4927 | -34.1043 | -9.0603
4.4988 | -34.1045 | -9.071
4.5049 | -34.1048 | -9.0816
4.511 | -34.1051 | -9.092
4.5172 | -34.1055 | -9.1032
4.5233 | -34.1057 | -9.115
4.5294 | -34.106 | -9.1258
4.5355 | -34.1062 | -9.1365
4.5416 | -34.1065 | -9.147
4.5477 | -34.1068 | -9.1577
4.5538 | -34.107 | -9.1679
4.5599 | -34.1074 | -9.1787
4.566 | -34.1076 | -9.1898
4.5721 | -34.1079 | -9.2004
4.5782 | -34.1083 | -9.2116
4.5843 | -34.1085 | -9.2225
4.5904 | -34.1089 | -9.2332
4.5965 | -34.1092 | -9.2434
4.6026 | -34.1095 | -9.2553
4.6087 | -34.1098 | -9.2657
4.6148 | -34.1101 | -9.2776
4.6209 | -34.1105 | -9.2875
4.627 | -34.1109 | -9.2989
4.6331 | -34.111 | -9.3092
4.6392 | -34.1113 | -9.3198
4.6453 | -34.1116 | -9.3306
4.6514 | -34.1119 | -9.3414
4.6576 | -34.112 | -9.3532
4.6637 | -34.1123 | -9.3624
4.6698 | -34.1126 | -9.3724
4.6759 | -34.1128 | -9.3836
4.682 | -34.113 | -9.3941
4.6881 | -34.1133 | -9.4043
4.6942 | -34.1136 | -9.415
4.7003 | -34.1138 | -9.4258
4.7064 | -34.1141 | -9.4365
4.7125 | -34.1145 | -9.4473
4.7186 | -34.1147 | -9.4572
4.7247 | -34.1151 | -9.4681
4.7308 | -34.1155 | -9.4788
4.7369 | -34.1157 | -9.4898
4.743 | -34.116 | -9.5014
4.7491 | -34.1162 | -9.5124
4.7552 | -34.1164 | -9.523
4.7613 | -34.1167 | -9.5339
4.7674 | -34.117 | -9.5439
4.7735 | -34.1172 | -9.5553
4.7796 | -34.1175 | -9.5664
4.7857 | -34.1176 | -9.5775
4.7918 | -34.1179 | -9.5875
4.7979 | -34.1183 | -9.598
4.8041 | -34.1185 | -9.6086
4.8102 | -34.1189 | -9.6191
4.8163 | -34.1192 | -9.6299
4.8224 | -34.1193 | -9.64
4.8285 | -34.1197 | -9.6505
4.8346 | -34.12 | -9.6613
4.8407 | -34.1203 | -9.672
4.8468 | -34.1206 | -9.6831
4.8529 | -34.1209 | -9.6942
4.859 | -34.1211 | -9.7051
4.8651 | -34.1215 | -9.7158
4.8712 | -34.1217 | -9.7266
4.8773 | -34.122 | -9.7375
4.8834 | -34.1223 | -9.7486
4.8895 | -34.1227 | -9.7599
4.8956 | -34.1229 | -9.7712
4.9017 | -34.1231 | -9.782
4.9078 | -34.1234 | -9.7932
4.9139 | -34.1237 | -9.8032
4.92 | -34.124 | -9.8136
4.9261 | -34.1244 | -9.8238
4.9322 | -34.1248 | -9.8344
4.9383 | -34.1252 | -9.8448
4.9445 | -34.1255 | -9.8562
4.9506 | -34.1258 | -9.8665
4.9567 | -34.1261 | -9.8772
4.9628 | -34.1264 | -9.8879
4.9689 | -34.1266 | -9.8991
4.975 | -34.1269 | -9.9098
4.9811 | -34.1273 | -9.9209
4.9872 | -34.1276 | -9.932
4.9933 | -34.128 | -9.9431
4.9994 | -34.1284 | -9.9542
5.0055 | -34.1288 | -9.9652
5.0116 | -34.1292 | -9.9767
5.0177 | -34.1296 | -9.9881
5.0238 | -34.13 | -9.9998
5.0299 | -34.1305 | -10.012
5.036 | -34.1308 | -10.0229
5.0421 | -34.131 | -10.034
5.0482 | -34.1314 | -10.0456
5.0543 | -34.1318 | -10.0563
5.0604 | -34.1322 | -10.0678
5.0665 | -34.1326 | -10.0788
5.0726 | -34.1329 | -10.0905
5.0787 | -34.1333 | -10.102
5.0848 | -34.1337 | -10.1134
5.091 | -34.134 | -10.1245
5.0971 | -34.1344 | -10.1359
5.1032 | -34.1348 | -10.1476
5.1093 | -34.1352 | -10.1588
5.1154 | -34.1355 | -10.1702
5.1215 | -34.1358 | -10.1815
5.1276 | -34.1362 | -10.1926
5.1337 | -34.1367 | -10.2037
5.1398 | -34.137 | -10.2148
5.1459 | -34.1374 | -10.2265
5.152 | -34.1378 | -10.2387
5.1581 | -34.1381 | -10.2507
5.1642 | -34.1384 | -10.2619
5.1703 | -34.1389 | -10.2733
5.1764 | -34.1392 | -10.2845
5.1825 | -34.1396 | -10.2961
5.1886 | -34.1399 | -10.3072
5.1947 | -34.1403 | -10.3192
5.2008 | -34.1408 | -10.3304
5.2069 | -34.1412 | -10.3428
5.213 | -34.1415 | -10.3537
5.2191 | -34.142 | -10.3653
5.2252 | -34.1423 | -10.3769
5.2314 | -34.1426 | -10.3889
5.2375 | -34.1431 | -10.4001
5.2436 | -34.1434 | -10.4116
5.2497 | -34.1439 | -10.4228
5.2558 | -34.1443 | -10.435
5.2619 | -34.1446 | -10.4471
5.268 | -34.145 | -10.4588
5.2741 | -34.1455 | -10.4712
5.2802 | -34.1459 | -10.4833
5.2863 | -34.1463 | -10.4948
5.2924 | -34.1466 | -10.5072
5.2985 | -34.1471 | -10.5197
5.3046 | -34.1476 | -10.5321
5.3107 | -34.148 | -10.5442
5.3168 | -34.1485 | -10.5559
5.3229 | -34.1489 | -10.5671
5.329 | -34.1493 | -10.5783
5.3351 | -34.1498 | -10.5898
5.3412 | -34.1503 | -10.6013
5.3473 | -34.1507 | -10.6134
5.3534 | -34.1512 | -10.6255
5.3595 | -34.1515 | -10.6374
5.3656 | -34.152 | -10.6496
5.3717 | -34.1524 | -10.6614
5.3779 | -34.1528 | -10.6735
5.384 | -34.1531 | -10.6855
5.3901 | -34.1536 | -10.6975
5.3962 | -34.1539 | -10.7095
5.4023 | -34.1544 | -10.7215
5.4084 | -34.1548 | -10.7331
5.4145 | -34.1552 | -10.7454
5.4206 | -34.1555 | -10.7578
5.4267 | -34.156 | -10.7703
5.4328 | -34.1564 | -10.7818
5.4389 | -34.1569 | -10.7933
5.445 | -34.1572 | -10.8058
5.4511 | -34.1577 | -10.8175
5.4572 | -34.1581 | -10.8298
5.4633 | -34.1585 | -10.8415
5.4694 | -34.1589 | -10.8532
5.4755 | -34.1593 | -10.8653
5.4816 | -34.1597 | -10.8771
5.4877 | -34.1601 | -10.889
5.4938 | -34.1605 | -10.9007
5.4999 | -34.1609 | -10.9126
5.506 | -34.1614 | -10.9246
5.5121 | -34.1618 | -10.9369
5.5183 | -34.1622 | -10.9482
5.5244 | -34.1625 | -10.9603
5.5305 | -34.1629 | -10.9724
5.5366 | -34.1633 | -10.9841
5.5427 | -34.1637 | -10.9965
5.5488 | -34.1641 | -11.0088
5.5549 | -34.1644 | -11.0206
5.561 | -34.1649 | -11.0331
5.5671 | -34.1653 | -11.0447
5.5732 | -34.1657 | -11.0569
5.5793 | -34.1661 | -11.0696
5.5854 | -34.1666 | -11.0817
5.5915 | -34.1669 | -11.0938
5.5976 | -34.1673 | -11.106
5.6037 | -34.1677 | -11.1179
5.6098 | -34.168 | -11.1297
5.6159 | -34.1685 | -11.1418
5.622 | -34.1689 | -11.1535
5.6281 | -34.1693 | -11.1655
5.6342 | -34.1697 | -11.1776
5.6403 | -34.1701 | -11.1891
5.6464 | -34.1705 | -11.2011
5.6525 | -34.1708 | -11.2129
5.6586 | -34.1713 | -11.2241
5.6648 | -34.1717 | -11.2366
5.6709 | -34.1721 | -11.2489
5.677 | -34.1725 | -11.2615
5.6831 | -34.1729 | -11.273
5.6892 | -34.1733 | -11.285
5.6953 | -34.1737 | -11.2968
5.7014 | -34.1741 | -11.3093
5.7075 | -34.1745 | -11.3215
5.7136 | -34.1749 | -11.3332
5.7197 | -34.1752 | -11.345
5.7258 | -34.1756 | -11.3568
5.7319 | -34.176 | -11.3687
5.738 | -34.1764 | -11.3806
5.7441 | -34.1768 | -11.3925
5.7502 | -34.1773 | -11.4041
5.7563 | -34.1777 | -11.4158
5.7624 | -34.1781 | -11.4283
5.7685 | -34.1784 | -11.4401
5.7746 | -34.1787 | -11.4518
5.7807 | -34.1791 | -11.4634
5.7868 | -34.1794 | -11.4753
5.7929 | -34.1798 | -11.4871
5.799 | -34.1802 | -11.4988
5.8052 | -34.1806 | -11.5108
5.8113 | -34.1809 | -11.5227
5.8174 | -34.1813 | -11.5347
5.8235 | -34.1816 | -11.5464
5.8296 | -34.182 | -11.5588
5.8357 | -34.1823 | -11.5706
5.8418 | -34.1827 | -11.5827
5.8479 | -34.183 | -11.5948
5.854 | -34.1834 | -11.6069
5.8601 | -34.1838 | -11.6186
5.8662 | -34.1841 | -11.6307
5.8723 | -34.1845 | -11.642
5.8784 | -34.1849 | -11.6538
5.8845 | -34.1853 | -11.6658
5.8906 | -34.1856 | -11.6775
5.8967 | -34.186 | -11.689
5.9028 | -34.1864 | -11.7009
5.9089 | -34.1867 | -11.7121
5.915 | -34.1871 | -11.7239
5.9211 | -34.1875 | -11.7358
5.9272 | -34.1879 | -11.7475
5.9333 | -34.1882 | -11.7587
5.9394 | -34.1887 | -11.7703
5.9455 | -34.189 | -11.7817
5.9517 | -34.1895 | -11.7934
5.9578 | -34.1899 | -11.8055
5.9639 | -34.1903 | -11.8169
5.97 | -34.1907 | -11.8288
5.9761 | -34.1911 | -11.8402
5.9822 | -34.1915 | -11.8521
5.9883 | -34.1919 | -11.8636
5.9944 | -34.1923 | -11.8753
6.0005 | -34.1927 | -11.8873
6.0066 | -34.1931 | -11.8986
6.0127 | -34.1935 | -11.9097
6.0188 | -34.1938 | -11.9217
6.0249 | -34.1942 | -11.9334
6.031 | -34.1946 | -11.9449
6.0371 | -34.1951 | -11.9568
6.0432 | -34.1955 | -11.9685
6.0493 | -34.1959 | -11.9799
6.0554 | -34.1964 | -11.9915
6.0615 | -34.1967 | -12.0036
6.0676 | -34.1971 | -12.0158
6.0737 | -34.1975 | -12.0276
6.0798 | -34.1979 | -12.0391
6.0859 | -34.1983 | -12.0507
6.0921 | -34.1987 | -12.0622
6.0982 | -34.1992 | -12.074
6.1043 | -34.1996 | -12.0856
6.1104 | -34.2 | -12.0977
6.1165 | -34.2004 | -12.1086
6.1226 | -34.2009 | -12.1197
6.1287 | -34.2013 | -12.1314
6.1348 | -34.2018 | -12.1431
6.1409 | -34.2021 | -12.1546
6.147 | -34.2026 | -12.1661
6.1531 | -34.203 | -12.1776
6.1592 | -34.2034 | -12.1892
6.1653 | -34.2038 | -12.2006
6.1714 | -34.2042 | -12.2122
6.1775 | -34.2046 | -12.2241
6.1836 | -34.205 | -12.2355
6.1897 | -34.2054 | -12.2471
6.1958 | -34.2058 | -12.2589
6.2019 | -34.2062 | -12.2707
6.208 | -34.2067 | -12.2826
6.2141 | -34.2071 | -12.2944
6.2202 | -34.2075 | -12.3064
6.2263 | -34.208 | -12.3178
6.2325 | -34.2084 | -12.3289
6.2386 | -34.2089 | -12.3402
6.2447 | -34.2092 | -12.352
6.2508 | -34.2096 | -12.3635
6.2569 | -34.2101 | -12.3753
6.263 | -34.2105 | -12.3866
6.2691 | -34.211 | -12.398
6.2752 | -34.2114 | -12.4095
6.2813 | -34.2118 | -12.421
6.2874 | -34.2122 | -12.4325
6.2935 | -34.2127 | -12.4439
6.2996 | -34.2131 | -12.4555
6.3057 | -34.2136 | -12.467
6.3118 | -34.214 | -12.4784
6.3179 | -34.2144 | -12.49
6.324 | -34.2148 | -12.5017
6.3301 | -34.2153 | -12.513
6.3362 | -34.2158 | -12.5246
6.3423 | -34.2162 | -12.5362
6.3484 | -34.2167 | -12.5479
6.3545 | -34.2172 | -12.5589
6.3606 | -34.2176 | -12.5702
6.3667 | -34.2181 | -12.5814
6.3728 | -34.2186 | -12.593
6.379 | -34.219 | -12.6045
6.3851 | -34.2194 | -12.616
6.3912 | -34.2199 | -12.6273
6.3973 | -34.2204 | -12.6385
6.4034 | -34.2208 | -12.6499
6.4095 | -34.2212 | -12.6616
6.4156 | -34.2216 | -12.6728
6.4217 | -34.222 | -12.6843
6.4278 | -34.2225 | -12.6957
6.4339 | -34.223 | -12.707
6.44 | -34.2234 | -12.7182
6.4461 | -34.2238 | -12.7301
6.4522 | -34.2243 | -12.741
6.4583 | -34.2247 | -12.7525
6.4644 | -34.225 | -12.7643
6.4705 | -34.2256 | -12.7759
6.4766 | -34.226 | -12.7877
6.4827 | -34.2265 | -12.7995
6.4888 | -34.2268 | -12.8108
6.4949 | -34.2273 | -12.8219
6.501 | -34.2277 | -12.8336
6.5071 | -34.2282 | -12.8452
6.5132 | -34.2287 | -12.8565
6.5194 | -34.2291 | -12.8678
6.5255 | -34.2296 | -12.8789
6.5316 | -34.23 | -12.8905
6.5377 | -34.2305 | -12.9022
6.5438 | -34.2309 | -12.9135
6.5499 | -34.2314 | -12.9248
6.556 | -34.2318 | -12.936
6.5621 | -34.2322 | -12.9476
6.5682 | -34.2326 | -12.9589
6.5743 | -34.2331 | -12.971
6.5804 | -34.2335 | -12.9824
6.5865 | -34.2338 | -12.9941
6.5926 | -34.2344 | -13.0055
6.5987 | -34.2348 | -13.0166
6.6048 | -34.2352 | -13.0287
6.6109 | -34.2357 | -13.04
6.617 | -34.2361 | -13.0522
6.6231 | -34.2365 | -13.0639
6.6292 | -34.237 | -13.0752
6.6353 | -34.2375 | -13.0876
6.6414 | -34.238 | -13.0993
6.6475 | -34.2384 | -13.1109
6.6536 | -34.2389 | -13.122
6.6597 | -34.2394 | -13.1332
6.6659 | -34.2398 | -13.1445
6.672 | -34.2403 | -13.156
6.6781 | -34.2407 | -13.1673
6.6842 | -34.2412 | -13.1789
6.6903 | -34.2417 | -13.1903
6.6964 | -34.2421 | -13.2018
6.7025 | -34.2425 | -13.2132
6.7086 | -34.2429 | -13.2245
6.7147 | -34.2434 | -13.236
6.7208 | -34.2438 | -13.2479
6.7269 | -34.2443 | -13.259
6.733 | -34.2446 | -13.2698
6.7391 | -34.2452 | -13.2814
6.7452 | -34.2457 | -13.2923
6.7513 | -34.2461 | -13.3034
6.7574 | -34.2466 | -13.3151
6.7635 | -34.2471 | -13.3261
6.7696 | -34.2476 | -13.3375
6.7757 | -34.2481 | -13.3492
6.7818 | -34.2485 | -13.3607
6.7879 | -34.249 | -13.3722
6.794 | -34.2494 | -13.3834
6.8001 | -34.2499 | -13.3947
6.8063 | -34.2503 | -13.4062
6.8124 | -34.2508 | -13.4177
6.8185 | -34.2512 | -13.4291
6.8246 | -34.2516 | -13.4405
6.8307 | -34.2521 | -13.4518
6.8368 | -34.2526 | -13.4634
6.8429 | -34.2531 | -13.4748
6.849 | -34.2535 | -13.4861
6.8551 | -34.254 | -13.4975
6.8612 | -34.2544 | -13.5088
6.8673 | -34.2549 | -13.52
6.8734 | -34.2553 | -13.532
6.8795 | -34.2558 | -13.5437
6.8856 | -34.2563 | -13.5546
6.8917 | -34.2567 | -13.5655
6.8978 | -34.2571 | -13.5768
6.9039 | -34.2576 | -13.588
6.91 | -34.2581 | -13.5994
6.9161 | -34.2585 | -13.6106
6.9222 | -34.2589 | -13.622
6.9283 | -34.2594 | -13.633
6.9344 | -34.2598 | -13.6442
6.9405 | -34.2603 | -13.6554
6.9466 | -34.2607 | -13.6669
6.9528 | -34.2612 | -13.6783
6.9589 | -34.2616 | -13.6895
6.965 | -34.2621 | -13.7008
6.9711 | -34.2626 | -13.7121
6.9772 | -34.2631 | -13.7233
6.9833 | -34.2635 | -13.7348
6.9894 | -34.2639 | -13.7465
6.9955 | -34.2644 | -13.7577
7.0016 | -34.2648 | -13.7694
7.0077 | -34.2653 | -13.781
7.0138 | -34.2657 | -13.7926
7.0199 | -34.2661 | -13.8037
7.026 | -34.2665 | -13.8149
7.0321 | -34.267 | -13.8263
7.0382 | -34.2674 | -13.8373
7.0443 | -34.2679 | -13.8489
7.0504 | -34.2683 | -13.8605
7.0565 | -34.2687 | -13.872
7.0626 | -34.2692 | -13.8832
7.0687 | -34.2696 | -13.8941
7.0748 | -34.2701 | -13.9052
7.0809 | -34.2705 | -13.9166
7.087 | -34.2709 | -13.928
7.0932 | -34.2713 | -13.9392
7.0993 | -34.2718 | -13.9506
7.1054 | -34.2722 | -13.9616
7.1115 | -34.2726 | -13.9728
7.1176 | -34.273 | -13.9839
7.1237 | -34.2734 | -13.9952
7.1298 | -34.2738 | -14.006
7.1359 | -34.2742 | -14.017
7.142 | -34.2746 | -14.0283
7.1481 | -34.275 | -14.0392
7.1542 | -34.2754 | -14.0503
7.1603 | -34.2758 | -14.0614
7.1664 | -34.2763 | -14.0727
7.1725 | -34.2768 | -14.0837
7.1786 | -34.2772 | -14.095
7.1847 | -34.2776 | -14.1061
7.1908 | -34.2781 | -14.1174
7.1969 | -34.2785 | -14.1285
7.203 | -34.2789 | -14.1398
7.2091 | -34.2793 | -14.1507
7.2152 | -34.2797 | -14.1619
7.2213 | -34.2802 | -14.1728
7.2274 | -34.2806 | -14.184
7.2335 | -34.281 | -14.1952
7.2397 | -34.2814 | -14.2063
7.2458 | -34.2818 | -14.217
7.2519 | -34.2823 | -14.2281
7.258 | -34.2827 | -14.2395
7.2641 | -34.2831 | -14.2505
7.2702 | -34.2836 | -14.2619
7.2763 | -34.284 | -14.2728
7.2824 | -34.2845 | -14.2838
7.2885 | -34.285 | -14.2949
7.2946 | -34.2854 | -14.3062
7.3007 | -34.2857 | -14.3173
7.3068 | -34.2861 | -14.3287
7.3129 | -34.2866 | -14.3396
7.319 | -34.287 | -14.3508
7.3251 | -34.2875 | -14.3615
7.3312 | -34.2879 | -14.3723
7.3373 | -34.2884 | -14.3832
7.3434 | -34.2889 | -14.394
7.3495 | -34.2893 | -14.4046
7.3556 | -34.2897 | -14.416
7.3617 | -34.2901 | -14.4275
7.3678 | -34.2906 | -14.4387
7.3739 | -34.291 | -14.4493
7.3801 | -34.2914 | -14.4603
7.3862 | -34.2918 | -14.4713
7.3923 | -34.2922 | -14.4823
7.3984 | -34.2926 | -14.4932
7.4045 | -34.2931 | -14.5042
7.4106 | -34.2935 | -14.5151
7.4167 | -34.2939 | -14.5263
7.4228 | -34.2944 | -14.5372
7.4289 | -34.2948 | -14.548
7.435 | -34.2953 | -14.5587
7.4411 | -34.2957 | -14.5695
7.4472 | -34.2961 | -14.5806
7.4533 | -34.2966 | -14.5915
7.4594 | -34.297 | -14.6028
7.4655 | -34.2975 | -14.6137
7.4716 | -34.2979 | -14.6249
7.4777 | -34.2983 | -14.6354
7.4838 | -34.2988 | -14.6466
7.4899 | -34.2992 | -14.6575
7.496 | -34.2997 | -14.6685
7.5021 | -34.3001 | -14.6794
7.5082 | -34.3005 | -14.6904
7.5143 | -34.301 | -14.7013
7.5204 | -34.3014 | -14.712
7.5266 | -34.3018 | -14.7229
7.5327 | -34.3022 | -14.7338
7.5388 | -34.3026 | -14.7446
7.5449 | -34.303 | -14.7558
7.551 | -34.3034 | -14.7669
7.5571 | -34.3039 | -14.7778
7.5632 | -34.3043 | -14.7886
7.5693 | -34.3048 | -14.7997
7.5754 | -34.3052 | -14.8106
7.5815 | -34.3057 | -14.8212
7.5876 | -34.306 | -14.8321
7.5937 | -34.3065 | -14.8429
7.5998 | -34.3069 | -14.8539
7.6059 | -34.3073 | -14.8648
7.612 | -34.3077 | -14.8757
7.6181 | -34.3081 | -14.8866
7.6242 | -34.3086 | -14.8975
7.6303 | -34.3091 | -14.9082
7.6364 | -34.3095 | -14.9193
7.6425 | -34.3099 | -14.9305
7.6486 | -34.3103 | -14.9415
7.6547 | -34.3108 | -14.9525
7.6608 | -34.3112 | -14.9634
7.667 | -34.3116 | -14.9738
7.6731 | -34.312 | -14.9848
7.6792 | -34.3124 | -14.9959
7.6853 | -34.3128 | -15.0066
7.6914 | -34.3132 | -15.0177
7.6975 | -34.3136 | -15.0287
7.7036 | -34.314 | -15.0397
7.7097 | -34.3145 | -15.0505
7.7158 | -34.3149 | -15.0613
7.7219 | -34.3153 | -15.0721
7.728 | -34.3157 | -15.0831
7.7341 | -34.3161 | -15.0941
7.7402 | -34.3166 | -15.1048
7.7463 | -34.317 | -15.1158
7.7524 | -34.3173 | -15.1269
7.7585 | -34.3178 | -15.138
7.7646 | -34.3182 | -15.1486
7.7707 | -34.3185 | -15.1593
7.7768 | -34.3189 | -15.1701
7.7829 | -34.3192 | -15.1811
7.789 | -34.3197 | -15.1917
7.7951 | -34.3201 | -15.203
7.8012 | -34.3205 | -15.2142
7.8073 | -34.3209 | -15.2249
7.8135 | -34.3213 | -15.2359
7.8196 | -34.3217 | -15.2467
7.8257 | -34.3222 | -15.258
7.8318 | -34.3226 | -15.2692
7.8379 | -34.323 | -15.2797
7.844 | -34.3233 | -15.2902
7.8501 | -34.3237 | -15.3013
7.8562 | -34.3241 | -15.3128
7.8623 | -34.3246 | -15.3237
7.8684 | -34.3249 | -15.3343
7.8745 | -34.3253 | -15.3451
7.8806 | -34.3257 | -15.3555
7.8867 | -34.3261 | -15.3662
7.8928 | -34.3266 | -15.3769
7.8989 | -34.327 | -15.3874
7.905 | -34.3275 | -15.3986
7.9111 | -34.3279 | -15.4095
7.9172 | -34.3283 | -15.4202
7.9233 | -34.3286 | -15.4311
7.9294 | -34.3291 | -15.4421
7.9355 | -34.3295 | -15.4525
7.9416 | -34.3298 | -15.4632
7.9477 | -34.3303 | -15.4737
7.9539 | -34.3307 | -15.4842
7.96 | -34.3311 | -15.4953
7.9661 | -34.3315 | -15.5062
7.9722 | -34.3319 | -15.5171
7.9783 | -34.3323 | -15.5279
7.9844 | -34.3328 | -15.5391
7.9905 | -34.3332 | -15.5499
7.9966 | -34.3335 | -15.5607
8.0027 | -34.3339 | -15.5718
8.0088 | -34.3343 | -15.5823
8.0149 | -34.3347 | -15.5929
8.021 | -34.3351 | -15.6038
8.0271 | -34.3356 | -15.615
8.0332 | -34.3359 | -15.6258
8.0393 | -34.3364 | -15.6368
8.0454 | -34.3367 | -15.6478
8.0515 | -34.3372 | -15.6585
8.0576 | -34.3375 | -15.6696
8.0637 | -34.3379 | -15.68
8.0698 | -34.3383 | -15.6904
8.0759 | -34.3387 | -15.701
8.082 | -34.3392 | -15.712
8.0881 | -34.3396 | -15.7227
8.0942 | -34.34 | -15.7339
8.1004 | -34.3403 | -15.745
8.1065 | -34.3408 | -15.7558
8.1126 | -34.3412 | -15.7665
8.1187 | -34.3416 | -15.7775
8.1248 | -34.342 | -15.7885
8.1309 | -34.3424 | -15.7993
8.137 | -34.3428 | -15.81
8.1431 | -34.3433 | -15.821
8.1492 | -34.3437 | -15.8319
8.1553 | -34.3441 | -15.8427
8.1614 | -34.3445 | -15.8538
8.1675 | -34.3449 | -15.8645
8.1736 | -34.3453 | -15.8756
8.1797 | -34.3456 | -15.8865
8.1858 | -34.346 | -15.8973
8.1919 | -34.3464 | -15.908
8.198 | -34.3468 | -15.9185
8.2041 | -34.3473 | -15.9297
8.2102 | -34.3477 | -15.9405
8.2163 | -34.3481 | -15.9513
8.2224 | -34.3485 | -15.9623
8.2285 | -34.3489 | -15.9729
8.2346 | -34.3494 | -15.9837
8.2408 | -34.3498 | -15.9946
8.2469 | -34.3502 | -16.0053
8.253 | -34.3505 | -16.0161
8.2591 | -34.351 | -16.0267
8.2652 | -34.3514 | -16.0373
8.2713 | -34.3519 | -16.0482
8.2774 | -34.3523 | -16.0591
8.2835 | -34.3527 | -16.0702
8.2896 | -34.3531 | -16.0806
8.2957 | -34.3535 | -16.0915
8.3018 | -34.3539 | -16.1022
8.3079 | -34.3544 | -16.113
8.314 | -34.3548 | -16.1236
8.3201 | -34.3552 | -16.1344
8.3262 | -34.3556 | -16.1451
8.3323 | -34.356 | -16.1559
8.3384 | -34.3565 | -16.1671
8.3445 | -34.3569 | -16.1775
8.3506 | -34.3573 | -16.1885
8.3567 | -34.3577 | -16.1995
8.3628 | -34.3581 | -16.2102
8.3689 | -34.3585 | -16.221
8.375 | -34.3589 | -16.2319
8.3812 | -34.3593 | -16.2426
8.3873 | -34.3597 | -16.2535
8.3934 | -34.3602 | -16.2643
8.3995 | -34.3606 | -16.2752
8.4056 | -34.361 | -16.286
8.4117 | -34.3614 | -16.2971
8.4178 | -34.3618 | -16.3075
8.4239 | -34.3622 | -16.3181
8.43 | -34.3627 | -16.3289
8.4361 | -34.3631 | -16.3401
8.4422 | -34.3635 | -16.3508
8.4483 | -34.3639 | -16.3616
8.4544 | -34.3643 | -16.3726
8.4605 | -34.3647 | -16.3833
8.4666 | -34.3651 | -16.3943
8.4727 | -34.3655 | -16.405
8.4788 | -34.3659 | -16.4156
8.4849 | -34.3663 | -16.4267
8.491 | -34.3667 | -16.4374
8.4971 | -34.3672 | -16.4482
8.5032 | -34.3676 | -16.4591
8.5093 | -34.368 | -16.4699
8.5154 | -34.3684 | -16.4804
8.5215 | -34.3688 | -16.4913
8.5277 | -34.3692 | -16.5021
8.5338 | -34.3696 | -16.5127
8.5399 | -34.37 | -16.5237
8.546 | -34.3705 | -16.5342
8.5521 | -34.3709 | -16.5451
8.5582 | -34.3713 | -16.556
8.5643 | -34.3717 | -16.5666
8.5704 | -34.3721 | -16.5771
8.5765 | -34.3725 | -16.5878
8.5826 | -34.373 | -16.5985
8.5887 | -34.3734 | -16.6092
8.5948 | -34.3738 | -16.6196
8.6009 | -34.3743 | -16.6303
8.607 | -34.3747 | -16.6413
8.6131 | -34.375 | -16.6523
8.6192 | -34.3755 | -16.6627
8.6253 | -34.3759 | -16.6736
8.6314 | -34.3764 | -16.6844
8.6375 | -34.3767 | -16.6953
8.6436 | -34.3772 | -16.7061
8.6497 | -34.3777 | -16.7169
8.6558 | -34.3781 | -16.7275
8.6619 | -34.3785 | -16.7385
8.6681 | -34.3789 | -16.7491
8.6742 | -34.3793 | -16.7601
8.6803 | -34.3797 | -16.7711
8.6864 | -34.3802 | -16.7821
8.6925 | -34.3806 | -16.7928
8.6986 | -34.381 | -16.8039
8.7047 | -34.3815 | -16.8147
8.7108 | -34.3819 | -16.8256
8.7169 | -34.3823 | -16.8362
8.723 | -34.3828 | -16.8469
8.7291 | -34.3832 | -16.8576
8.7352 | -34.3837 | -16.8686
8.7413 | -34.3841 | -16.8793
8.7474 | -34.3845 | -16.8898
8.7535 | -34.385 | -16.9007
8.7596 | -34.3854 | -16.9119
8.7657 | -34.3859 | -16.9226
8.7718 | -34.3863 | -16.9332
8.7779 | -34.3867 | -16.9444
8.784 | -34.3871 | -16.955
8.7901 | -34.3875 | -16.9662
8.7962 | -34.3879 | -16.977
8.8023 | -34.3884 | -16.9875
8.8084 | -34.3888 | -16.9984
8.8146 | -34.3893 | -17.0091
8.8207 | -34.3897 | -17.0199
8.8268 | -34.3902 | -17.0306
8.8329 | -34.3906 | -17.0414
8.839 | -34.3911 | -17.0523
8.8451 | -34.3915 | -17.0629
8.8512 | -34.392 | -17.0737
8.8573 | -34.3924 | -17.0843
8.8634 | -34.3929 | -17.095
8.8695 | -34.3933 | -17.1059
8.8756 | -34.3937 | -17.1164
8.8817 | -34.3942 | -17.1274
8.8878 | -34.3946 | -17.1384
8.8939 | -34.395 | -17.1489
8.9 | -34.3954 | -17.1593
8.9061 | -34.3959 | -17.1701
8.9122 | -34.3963 | -17.1808
8.9183 | -34.3968 | -17.1914
8.9244 | -34.3972 | -17.202
8.9305 | -34.3977 | -17.2127
8.9366 | -34.3981 | -17.2234
8.9427 | -34.3986 | -17.2342
8.9488 | -34.3991 | -17.2449
8.955 | -34.3995 | -17.2556
8.9611 | -34.3999 | -17.2665
8.9672 | -34.4004 | -17.2776
8.9733 | -34.4009 | -17.2882
8.9794 | -34.4013 | -17.2991
8.9855 | -34.4018 | -17.3097
8.9916 | -34.4022 | -17.3203
8.9977 | -34.4027 | -17.331
9.0038 | -34.4032 | -17.3419
9.0099 | -34.4036 | -17.3527
9.016 | -34.4041 | -17.3634
9.0221 | -34.4046 | -17.3744
9.0282 | -34.405 | -17.3853
9.0343 | -34.4055 | -17.3962
9.0404 | -34.406 | -17.4068
9.0465 | -34.4065 | -17.4179
9.0526 | -34.4069 | -17.4286
9.0587 | -34.4074 | -17.4393
9.0648 | -34.4079 | -17.45
9.0709 | -34.4084 | -17.4605
9.077 | -34.4088 | -17.4717
9.0831 | -34.4092 | -17.4825
9.0892 | -34.4097 | -17.4932
9.0953 | -34.4101 | -17.5041
9.1015 | -34.4106 | -17.5148
9.1076 | -34.4111 | -17.5256
9.1137 | -34.4115 | -17.5363
9.1198 | -34.412 | -17.5467
9.1259 | -34.4125 | -17.5575
9.132 | -34.413 | -17.5685
9.1381 | -34.4135 | -17.5791
9.1442 | -34.414 | -17.5894
9.1503 | -34.4144 | -17.6003
9.1564 | -34.4149 | -17.611
9.1625 | -34.4154 | -17.6218
9.1686 | -34.4159 | -17.6326
9.1747 | -34.4163 | -17.6434
9.1808 | -34.4168 | -17.6543
9.1869 | -34.4173 | -17.6653
9.193 | -34.4177 | -17.6762
9.1991 | -34.4182 | -17.6867
9.2052 | -34.4187 | -17.6973
9.2113 | -34.4192 | -17.7083
9.2174 | -34.4197 | -17.7196
9.2235 | -34.4202 | -17.73
9.2296 | -34.4207 | -17.7409
9.2357 | -34.4211 | -17.7517
9.2419 | -34.4216 | -17.7624
9.248 | -34.4221 | -17.7733
9.2541 | -34.4226 | -17.7842
9.2602 | -34.423 | -17.795
9.2663 | -34.4235 | -17.8056
9.2724 | -34.424 | -17.8164
9.2785 | -34.4245 | -17.8272
9.2846 | -34.425 | -17.8379
9.2907 | -34.4254 | -17.8486
9.2968 | -34.4259 | -17.8595
9.3029 | -34.4264 | -17.8704
9.309 | -34.4269 | -17.881
9.3151 | -34.4273 | -17.8917
9.3212 | -34.4279 | -17.9025
9.3273 | -34.4284 | -17.9131
9.3334 | -34.4289 | -17.9237
9.3395 | -34.4294 | -17.9344
9.3456 | -34.43 | -17.9449
9.3517 | -34.4304 | -17.9553
9.3578 | -34.4309 | -17.9658
9.3639 | -34.4314 | -17.9767
9.37 | -34.4319 | -17.9875
9.3761 | -34.4324 | -17.9981
9.3822 | -34.4329 | -18.0087
9.3884 | -34.4334 | -18.0194
9.3945 | -34.4339 | -18.0302
9.4006 | -34.4344 | -18.0409
9.4067 | -34.4349 | -18.0518
9.4128 | -34.4354 | -18.0624
9.4189 | -34.4359 | -18.0733
9.425 | -34.4364 | -18.0838
9.4311 | -34.4369 | -18.0945
9.4372 | -34.4374 | -18.1052
9.4433 | -34.4379 | -18.116
9.4494 | -34.4384 | -18.1267
9.4555 | -34.4389 | -18.1374
9.4616 | -34.4394 | -18.148
9.4677 | -34.44 | -18.1585
9.4738 | -34.4404 | -18.1696
9.4799 | -34.4409 | -18.1802
9.486 | -34.4413 | -18.1912
9.4921 | -34.4419 | -18.2023
9.4982 | -34.4424 | -18.2125
9.5043 | -34.4429 | -18.223
9.5104 | -34.4434 | -18.2336
9.5165 | -34.4439 | -18.2442
9.5226 | -34.4444 | -18.2551
9.5288 | -34.4449 | -18.2657
9.5349 | -34.4454 | -18.2764
9.541 | -34.4459 | -18.2871
9.5471 | -34.4464 | -18.2979
9.5532 | -34.4469 | -18.3085
9.5593 | -34.4475 | -18.3192
9.5654 | -34.448 | -18.3299
9.5715 | -34.4485 | -18.3408
9.5776 | -34.449 | -18.3518
9.5837 | -34.4495 | -18.3622
9.5898 | -34.4499 | -18.3731
9.5959 | -34.4504 | -18.3837
9.602 | -34.4509 | -18.3944
9.6081 | -34.4514 | -18.4052
9.6142 | -34.452 | -18.4158
9.6203 | -34.4524 | -18.4265
9.6264 | -34.4529 | -18.4368
9.6325 | -34.4533 | -18.4477
9.6386 | -34.4539 | -18.4583
9.6447 | -34.4543 | -18.4689
9.6508 | -34.4549 | -18.4796
9.6569 | -34.4554 | -18.4903
9.663 | -34.4559 | -18.501
9.6691 | -34.4564 | -18.5117
9.6753 | -34.457 | -18.5223
9.6814 | -34.4575 | -18.5328
9.6875 | -34.4581 | -18.5432
9.6936 | -34.4586 | -18.5538
9.6997 | -34.4591 | -18.5646
9.7058 | -34.4596 | -18.5752
9.7119 | -34.4601 | -18.5858
9.718 | -34.4606 | -18.5967
9.7241 | -34.4611 | -18.6074
9.7302 | -34.4615 | -18.6182
9.7363 | -34.4621 | -18.6286
9.7424 | -34.4626 | -18.6391
9.7485 | -34.4631 | -18.6497
9.7546 | -34.4636 | -18.6601
9.7607 | -34.4641 | -18.6709
9.7668 | -34.4646 | -18.6819
9.7729 | -34.4652 | -18.6924
9.779 | -34.4658 | -18.7029
9.7851 | -34.4662 | -18.7137
9.7912 | -34.4667 | -18.7248
9.7973 | -34.4672 | -18.7357
9.8034 | -34.4677 | -18.7465
9.8095 | -34.4682 | -18.757
9.8157 | -34.4688 | -18.7676
9.8218 | -34.4693 | -18.778
9.8279 | -34.4698 | -18.7887
9.834 | -34.4703 | -18.7992
9.8401 | -34.4708 | -18.8098
9.8462 | -34.4714 | -18.8201
9.8523 | -34.4719 | -18.8306
9.8584 | -34.4724 | -18.8415
9.8645 | -34.4729 | -18.8521
9.8706 | -34.4734 | -18.8627
9.8767 | -34.4739 | -18.8729
9.8828 | -34.4745 | -18.8831
9.8889 | -34.475 | -18.894
9.895 | -34.4755 | -18.9043
9.9011 | -34.476 | -18.9144
9.9072 | -34.4765 | -18.9247
9.9133 | -34.477 | -18.9352
9.9194 | -34.4776 | -18.9455
9.9255 | -34.4781 | -18.9562
9.9316 | -34.4786 | -18.9665
9.9377 | -34.4792 | -18.9769
9.9438 | -34.4797 | -18.9874
9.9499 | -34.4802 | -18.9981
9.956 | -34.4808 | -19.0084
9.9622 | -34.4813 | -19.0189
9.9683 | -34.4818 | -19.0295
9.9744 | -34.4823 | -19.04
9.9805 | -34.4828 | -19.0505
9.9866 | -34.4833 | -19.0607
9.9927 | -34.4839 | -19.0714
9.9988 | -34.4844 | -19.0818
10.0049 | -34.4849 | -19.0923
10.011 | -34.4855 | -19.1028
10.0171 | -34.486 | -19.1133
10.0232 | -34.4865 | -19.1235
10.0293 | -34.4871 | -19.1338
10.0354 | -34.4877 | -19.144
10.0415 | -34.4882 | -19.1545
10.0476 | -34.4887 | -19.1653
10.0537 | -34.4892 | -19.176
10.0598 | -34.4897 | -19.1866
10.0659 | -34.4902 | -19.1971
10.072 | -34.4908 | -19.2075
10.0781 | -34.4913 | -19.2179
10.0842 | -34.4918 | -19.2285
10.0903 | -34.4924 | -19.2393
10.0964 | -34.493 | -19.2494
10.1026 | -34.4935 | -19.2599
10.1087 | -34.4941 | -19.2704
10.1148 | -34.4946 | -19.281
10.1209 | -34.4951 | -19.2913
10.127 | -34.4956 | -19.3015
10.1331 | -34.4962 | -19.3118
10.1392 | -34.4967 | -19.3222
10.1453 | -34.4973 | -19.3327
10.1514 | -34.4978 | -19.3429
10.1575 | -34.4983 | -19.3533
10.1636 | -34.4989 | -19.363
10.1697 | -34.4994 | -19.3735
10.1758 | -34.4999 | -19.3838
10.1819 | -34.5005 | -19.3943
10.188 | -34.501 | -19.4045
10.1941 | -34.5015 | -19.415
10.2002 | -34.5019 | -19.4248
10.2063 | -34.5025 | -19.4355
10.2124 | -34.503 | -19.4459
10.2185 | -34.5035 | -19.4562
10.2246 | -34.5041 | -19.4666
10.2307 | -34.5046 | -19.4771
10.2368 | -34.5051 | -19.4867
10.2429 | -34.5057 | -19.4968
10.2491 | -34.5062 | -19.507
10.2552 | -34.5068 | -19.5177
10.2613 | -34.5073 | -19.5283
10.2674 | -34.5078 | -19.5388
10.2735 | -34.5083 | -19.5494
10.2796 | -34.5088 | -19.5598
10.2857 | -34.5093 | -19.5702
10.2918 | -34.5099 | -19.5807
10.2979 | -34.5104 | -19.5912
10.304 | -34.511 | -19.6017
10.3101 | -34.5115 | -19.6121
10.3162 | -34.5121 | -19.6224
10.3223 | -34.5126 | -19.6328
10.3284 | -34.5131 | -19.6431
10.3345 | -34.5137 | -19.6533
10.3406 | -34.5142 | -19.6634
10.3467 | -34.5147 | -19.6737
10.3528 | -34.5152 | -19.6839
10.3589 | -34.5157 | -19.6944
10.365 | -34.5163 | -19.7049
10.3711 | -34.5168 | -19.7155
10.3772 | -34.5172 | -19.7258
10.3833 | -34.5178 | -19.7361
10.3895 | -34.5183 | -19.7463
10.3956 | -34.5188 | -19.7564
10.4017 | -34.5193 | -19.7667
10.4078 | -34.5199 | -19.777
10.4139 | -34.5204 | -19.7873
10.42 | -34.521 | -19.7977
10.4261 | -34.5215 | -19.8079
10.4322 | -34.5221 | -19.8181
10.4383 | -34.5226 | -19.8284
10.4444 | -34.5231 | -19.8383
10.4505 | -34.5237 | -19.8486
10.4566 | -34.5242 | -19.8589
10.4627 | -34.5247 | -19.8691
10.4688 | -34.5253 | -19.879
10.4749 | -34.5258 | -19.8892
10.481 | -34.5263 | -19.8994
10.4871 | -34.5269 | -19.9098
10.4932 | -34.5274 | -19.92
10.4993 | -34.528 | -19.9302
10.5054 | -34.5285 | -19.9405
10.5115 | -34.5291 | -19.9508
10.5176 | -34.5296 | -19.9608
10.5237 | -34.5302 | -19.9709
10.5298 | -34.5307 | -19.9809
10.536 | -34.5313 | -19.9914
10.5421 | -34.5318 | -20.0012
10.5482 | -34.5323 | -20.0117
10.5543 | -34.5329 | -20.0219
10.5604 | -34.5334 | -20.0318
10.5665 | -34.534 | -20.0419
10.5726 | -34.5345 | -20.0519
10.5787 | -34.535 | -20.0623
10.5848 | -34.5356 | -20.0721
10.5909 | -34.5361 | -20.0824
10.597 | -34.5366 | -20.0926
10.6031 | -34.5372 | -20.1029
10.6092 | -34.5377 | -20.1134
10.6153 | -34.5382 | -20.1238
10.6214 | -34.5388 | -20.1339
10.6275 | -34.5394 | -20.1439
10.6336 | -34.5399 | -20.1545
10.6397 | -34.5404 | -20.1646
10.6458 | -34.5409 | -20.1745
10.6519 | -34.5414 | -20.1847
10.658 | -34.5419 | -20.1946
10.6641 | -34.5425 | -20.2048
10.6702 | -34.543 | -20.2152
10.6764 | -34.5436 | -20.2253
10.6825 | -34.5442 | -20.2356
10.6886 | -34.5447 | -20.2459
10.6947 | -34.5453 | -20.2558
10.7008 | -34.5458 | -20.2659
10.7069 | -34.5464 | -20.2761
10.713 | -34.5469 | -20.2862
10.7191 | -34.5475 | -20.2965
10.7252 | -34.548 | -20.3064
10.7313 | -34.5486 | -20.3166
10.7374 | -34.5491 | -20.3265
10.7435 | -34.5497 | -20.3369
10.7496 | -34.5502 | -20.3468
10.7557 | -34.5507 | -20.357
10.7618 | -34.5513 | -20.3671
10.7679 | -34.5518 | -20.3772
10.774 | -34.5523 | -20.3873
10.7801 | -34.5529 | -20.3978
10.7862 | -34.5534 | -20.4078
10.7923 | -34.554 | -20.418
10.7984 | -34.5545 | -20.428
10.8045 | -34.5551 | -20.4383
10.8106 | -34.5556 | -20.4484
10.8168 | -34.5562 | -20.4585
10.8229 | -34.5567 | -20.4687
10.829 | -34.5572 | -20.4789
10.8351 | -34.5578 | -20.489
10.8412 | -34.5583 | -20.499
10.8473 | -34.5588 | -20.5089
10.8534 | -34.5594 | -20.5192
10.8595 | -34.5599 | -20.5291
10.8656 | -34.5605 | -20.539
10.8717 | -34.561 | -20.5488
10.8778 | -34.5616 | -20.5588
10.8839 | -34.5621 | -20.5687
10.89 | -34.5626 | -20.5787
10.8961 | -34.5632 | -20.5888
10.9022 | -34.5638 | -20.5987
10.9083 | -34.5643 | -20.609
10.9144 | -34.5649 | -20.619
10.9205 | -34.5655 | -20.6291
10.9266 | -34.566 | -20.6394
10.9327 | -34.5666 | -20.6495
10.9388 | -34.5672 | -20.6595
10.9449 | -34.5677 | -20.6696
10.951 | -34.5682 | -20.6796
10.9571 | -34.5688 | -20.6898
10.9633 | -34.5693 | -20.7001
10.9694 | -34.5699 | -20.7105
10.9755 | -34.5704 | -20.7205
10.9816 | -34.571 | -20.7306
10.9877 | -34.5715 | -20.7407
10.9938 | -34.5721 | -20.751
10.9999 | -34.5726 | -20.7609
11.006 | -34.5732 | -20.7708
11.0121 | -34.5737 | -20.7812
11.0182 | -34.5743 | -20.7909
11.0243 | -34.5749 | -20.801
11.0304 | -34.5754 | -20.811
11.0365 | -34.5759 | -20.821
11.0426 | -34.5765 | -20.8308
11.0487 | -34.5771 | -20.8409
11.0548 | -34.5776 | -20.8509
11.0609 | -34.5782 | -20.8612
11.067 | -34.5787 | -20.8714
11.0731 | -34.5793 | -20.8815
11.0792 | -34.5798 | -20.8914
11.0853 | -34.5804 | -20.9012
11.0914 | -34.5809 | -20.9112
11.0975 | -34.5815 | -20.9214
11.1037 | -34.582 | -20.9312
11.1098 | -34.5826 | -20.9415
11.1159 | -34.5832 | -20.9516
11.122 | -34.5837 | -20.9618
11.1281 | -34.5842 | -20.9716
11.1342 | -34.5848 | -20.9814
11.1403 | -34.5853 | -20.9917
11.1464 | -34.5859 | -21.0017
11.1525 | -34.5864 | -21.0117
11.1586 | -34.587 | -21.0223
11.1647 | -34.5876 | -21.0321
11.1708 | -34.5881 | -21.042
11.1769 | -34.5886 | -21.0516
11.183 | -34.5892 | -21.0615
11.1891 | -34.5898 | -21.0718
11.1952 | -34.5903 | -21.0821
11.2013 | -34.5908 | -21.0921
11.2074 | -34.5914 | -21.1019
11.2135 | -34.5919 | -21.1119
11.2196 | -34.5924 | -21.1218
11.2257 | -34.5929 | -21.1317
11.2318 | -34.5934 | -21.1419
11.2379 | -34.5939 | -21.1518
11.244 | -34.5945 | -21.1617
11.2502 | -34.595 | -21.1717
11.2563 | -34.5956 | -21.1816
11.2624 | -34.5961 | -21.1916
11.2685 | -34.5967 | -21.2014
11.2746 | -34.5972 | -21.2111
11.2807 | -34.5977 | -21.2216
11.2868 | -34.5982 | -21.2318
11.2929 | -34.5988 | -21.2415
11.299 | -34.5994 | -21.2517
11.3051 | -34.5999 | -21.2616
11.3112 | -34.6004 | -21.2717
11.3173 | -34.601 | -21.2817
11.3234 | -34.6015 | -21.2917
11.3295 | -34.6021 | -21.3015
11.3356 | -34.6026 | -21.3114
11.3417 | -34.6032 | -21.3215
11.3478 | -34.6037 | -21.3315
11.3539 | -34.6042 | -21.3414
11.36 | -34.6047 | -21.3512
11.3661 | -34.6053 | -21.3611
11.3722 | -34.6059 | -21.3709
11.3783 | -34.6064 | -21.3809
11.3844 | -34.6069 | -21.3909
11.3906 | -34.6075 | -21.4011
11.3967 | -34.608 | -21.4108
11.4028 | -34.6085 | -21.421
11.4089 | -34.6091 | -21.4307
11.415 | -34.6096 | -21.4406
11.4211 | -34.6101 | -21.4504
11.4272 | -34.6107 | -21.4604
11.4333 | -34.6112 | -21.4704
11.4394 | -34.6118 | -21.4806
11.4455 | -34.6122 | -21.4906
11.4516 | -34.6128 | -21.5004
11.4577 | -34.6134 | -21.5106
11.4638 | -34.614 | -21.52
11.4699 | -34.6145 | -21.5303
11.476 | -34.615 | -21.5407
11.4821 | -34.6156 | -21.5503
11.4882 | -34.6161 | -21.5606
11.4943 | -34.6166 | -21.5704
11.5004 | -34.6172 | -21.5803
11.5065 | -34.6178 | -21.5902
11.5126 | -34.6183 | -21.5998
11.5187 | -34.6189 | -21.6097
11.5248 | -34.6194 | -21.6197
11.5309 | -34.62 | -21.6296
11.5371 | -34.6205 | -21.6394
11.5432 | -34.621 | -21.6495
11.5493 | -34.6215 | -21.6596
11.5554 | -34.6221 | -21.6693
11.5615 | -34.6226 | -21.6795
11.5676 | -34.6232 | -21.6896
11.5737 | -34.6238 | -21.6992
11.5798 | -34.6243 | -21.7089
11.5859 | -34.6249 | -21.7187
11.592 | -34.6254 | -21.7286
11.5981 | -34.626 | -21.7383
11.6042 | -34.6265 | -21.7483
11.6103 | -34.6271 | -21.7581
11.6164 | -34.6276 | -21.7678
11.6225 | -34.6281 | -21.7777
11.6286 | -34.6287 | -21.7877
11.6347 | -34.6292 | -21.7978
11.6408 | -34.6298 | -21.8076
11.6469 | -34.6304 | -21.8178
11.653 | -34.6309 | -21.8274
11.6591 | -34.6315 | -21.8373
11.6652 | -34.6321 | -21.8473
11.6713 | -34.6326 | -21.8572
11.6775 | -34.6332 | -21.8668
11.6836 | -34.6337 | -21.8768
11.6897 | -34.6343 | -21.8865
11.6958 | -34.6348 | -21.8967
11.7019 | -34.6354 | -21.9067
11.708 | -34.6359 | -21.9167
11.7141 | -34.6365 | -21.9266
11.7202 | -34.6371 | -21.9362
11.7263 | -34.6376 | -21.9462
11.7324 | -34.6382 | -21.9562
11.7385 | -34.6387 | -21.9662
11.7446 | -34.6392 | -21.9762
11.7507 | -34.6398 | -21.986
11.7568 | -34.6403 | -21.996
11.7629 | -34.6409 | -22.0058
11.769 | -34.6415 | -22.0157
11.7751 | -34.642 | -22.0253
11.7812 | -34.6426 | -22.0349
11.7873 | -34.6431 | -22.0451
11.7934 | -34.6437 | -22.0549
11.7995 | -34.6443 | -22.0645
11.8056 | -34.6448 | -22.0747
11.8117 | -34.6453 | -22.0846
11.8178 | -34.6459 | -22.0943
11.824 | -34.6465 | -22.104
11.8301 | -34.6471 | -22.1135
11.8362 | -34.6476 | -22.1231
11.8423 | -34.6481 | -22.133
11.8484 | -34.6487 | -22.1428
11.8545 | -34.6492 | -22.1528
11.8606 | -34.6498 | -22.1627
11.8667 | -34.6504 | -22.1726
11.8728 | -34.6509 | -22.1824
11.8789 | -34.6514 | -22.1926
11.885 | -34.6519 | -22.2023
11.8911 | -34.6526 | -22.2123
11.8972 | -34.6531 | -22.2218
11.9033 | -34.6536 | -22.2312
11.9094 | -34.6541 | -22.2407
11.9155 | -34.6546 | -22.2506
11.9216 | -34.6553 | -22.2602
11.9277 | -34.6558 | -22.2698
11.9338 | -34.6563 | -22.2797
11.9399 | -34.6569 | -22.2898
11.946 | -34.6575 | -22.2996
11.9521 | -34.658 | -22.3093
11.9582 | -34.6586 | -22.3189
11.9644 | -34.6592 | -22.3286
11.9705 | -34.6597 | -22.3386
11.9766 | -34.6603 | -22.3483
11.9827 | -34.6608 | -22.358
11.9888 | -34.6614 | -22.3676
11.9949 | -34.6619 | -22.3773
12.001 | -34.6625 | -22.3871
12.0071 | -34.6631 | -22.3967
12.0132 | -34.6636 | -22.4062
12.0193 | -34.6642 | -22.4162
12.0254 | -34.6647 | -22.4257
12.0315 | -34.6653 | -22.4353
12.0376 | -34.6658 | -22.4449
12.0437 | -34.6664 | -22.4544
12.0498 | -34.6669 | -22.4644
12.0559 | -34.6675 | -22.4742
12.062 | -34.6681 | -22.4838
12.0681 | -34.6686 | -22.4935
12.0742 | -34.6692 | -22.5034
12.0803 | -34.6697 | -22.513
12.0864 | -34.6703 | -22.5228
12.0925 | -34.6709 | -22.5325
12.0986 | -34.6715 | -22.5419
12.1047 | -34.6721 | -22.5516
12.1109 | -34.6726 | -22.5615
12.117 | -34.6732 | -22.5711
12.1231 | -34.6738 | -22.5808
12.1292 | -34.6743 | -22.5907
12.1353 | -34.6749 | -22.6001
12.1414 | -34.6755 | -22.61
12.1475 | -34.676 | -22.6197
12.1536 | -34.6766 | -22.6295
12.1597 | -34.6771 | -22.6393
12.1658 | -34.6777 | -22.649
12.1719 | -34.6783 | -22.6587
12.178 | -34.6789 | -22.6686
12.1841 | -34.6794 | -22.6784
12.1902 | -34.68 | -22.6881
12.1963 | -34.6806 | -22.6978
12.2024 | -34.6812 | -22.7073
12.2085 | -34.6817 | -22.717
12.2146 | -34.6823 | -22.7266
12.2207 | -34.6828 | -22.7363
12.2268 | -34.6834 | -22.746
12.2329 | -34.6839 | -22.7556
12.239 | -34.6845 | -22.765
12.2451 | -34.685 | -22.7745
12.2513 | -34.6857 | -22.7842
12.2574 | -34.6862 | -22.7941
12.2635 | -34.6868 | -22.8034
12.2696 | -34.6873 | -22.8129
12.2757 | -34.6879 | -22.8222
12.2818 | -34.6884 | -22.8317
12.2879 | -34.6889 | -22.8412
12.294 | -34.6895 | -22.8506
12.3001 | -34.6901 | -22.8604
12.3062 | -34.6907 | -22.87
12.3123 | -34.6913 | -22.8796
12.3184 | -34.6918 | -22.8891
12.3245 | -34.6924 | -22.8988
12.3306 | -34.693 | -22.9082
12.3367 | -34.6935 | -22.9176
12.3428 | -34.6941 | -22.927
12.3489 | -34.6947 | -22.9361
12.355 | -34.6953 | -22.9456
12.3611 | -34.6958 | -22.9552
12.3672 | -34.6964 | -22.9647
12.3733 | -34.6969 | -22.9746
12.3794 | -34.6975 | -22.9836
12.3855 | -34.698 | -22.9931
12.3916 | -34.6986 | -23.0028
12.3978 | -34.6992 | -23.0126
12.4039 | -34.6997 | -23.0222
12.41 | -34.7003 | -23.0319
12.4161 | -34.7009 | -23.0412
12.4222 | -34.7015 | -23.0506
12.4283 | -34.702 | -23.0602
12.4344 | -34.7026 | -23.0697
12.4405 | -34.7032 | -23.0793
12.4466 | -34.7037 | -23.0887
12.4527 | -34.7043 | -23.0983
12.4588 | -34.7049 | -23.1078
12.4649 | -34.7054 | -23.1176
12.471 | -34.706 | -23.1275
12.4771 | -34.7066 | -23.1365
12.4832 | -34.7071 | -23.146
12.4893 | -34.7077 | -23.1553
12.4954 | -34.7083 | -23.1647
12.5015 | -34.7089 | -23.174
12.5076 | -34.7095 | -23.1836
12.5137 | -34.71 | -23.1935
12.5198 | -34.7106 | -23.2026
12.5259 | -34.7112 | -23.2121
12.532 | -34.7118 | -23.2217
12.5382 | -34.7123 | -23.2311
12.5443 | -34.7129 | -23.2407
12.5504 | -34.7134 | -23.2502
12.5565 | -34.7141 | -23.2595
12.5626 | -34.7146 | -23.2692
12.5687 | -34.7152 | -23.2788
12.5748 | -34.7158 | -23.2883
12.5809 | -34.7163 | -23.2974
12.587 | -34.7168 | -23.3072
12.5931 | -34.7175 | -23.3166
12.5992 | -34.718 | -23.3263
12.6053 | -34.7186 | -23.3357
12.6114 | -34.7192 | -23.3452
12.6175 | -34.7197 | -23.3548
12.6236 | -34.7203 | -23.364
12.6297 | -34.7209 | -23.3733
12.6358 | -34.7215 | -23.3825
12.6419 | -34.7221 | -23.3922
12.648 | -34.7227 | -23.4019
12.6541 | -34.7232 | -23.4113
12.6602 | -34.7239 | -23.4202
12.6663 | -34.7244 | -23.4296
12.6724 | -34.725 | -23.4391
12.6785 | -34.7255 | -23.4487
12.6847 | -34.7261 | -23.4582
12.6908 | -34.7267 | -23.4676
12.6969 | -34.7273 | -23.4769
12.703 | -34.7279 | -23.4864
12.7091 | -34.7284 | -23.4959
12.7152 | -34.729 | -23.5053
12.7213 | -34.7296 | -23.5148
12.7274 | -34.7301 | -23.5243
12.7335 | -34.7307 | -23.5331
12.7396 | -34.7313 | -23.5422
12.7457 | -34.7319 | -23.5516
12.7518 | -34.7324 | -23.561
12.7579 | -34.733 | -23.5707
12.764 | -34.7335 | -23.5801
12.7701 | -34.734 | -23.5892
12.7762 | -34.7346 | -23.5986
12.7823 | -34.7352 | -23.6081
12.7884 | -34.7357 | -23.6174
12.7945 | -34.7363 | -23.6265
12.8006 | -34.7369 | -23.6352
12.8067 | -34.7375 | -23.6448
12.8128 | -34.738 | -23.6539
12.8189 | -34.7385 | -23.663
12.8251 | -34.739 | -23.6722
12.8312 | -34.7396 | -23.6815
12.8373 | -34.7402 | -23.6908
12.8434 | -34.7408 | -23.7001
12.8495 | -34.7413 | -23.7093
12.8556 | -34.7418 | -23.7186
12.8617 | -34.7424 | -23.7277
12.8678 | -34.743 | -23.7371
12.8739 | -34.7435 | -23.7459
12.88 | -34.7441 | -23.7554
12.8861 | -34.7446 | -23.7647
12.8922 | -34.7452 | -23.774
12.8983 | -34.7458 | -23.783
12.9044 | -34.7464 | -23.7923
12.9105 | -34.747 | -23.8014
12.9166 | -34.7476 | -23.811
12.9227 | -34.7481 | -23.8207
12.9288 | -34.7488 | -23.8301
12.9349 | -34.7493 | -23.8397
12.941 | -34.75 | -23.849
12.9471 | -34.7506 | -23.8582
12.9532 | -34.7512 | -23.8675
12.9593 | -34.7517 | -23.877
12.9654 | -34.7523 | -23.8865
12.9716 | -34.7528 | -23.8958
12.9777 | -34.7534 | -23.9047
12.9838 | -34.7539 | -23.9138
12.9899 | -34.7545 | -23.9231
12.996 | -34.755 | -23.9326
13.0021 | -34.7556 | -23.9419
13.0082 | -34.7561 | -23.9512
13.0143 | -34.7567 | -23.96
13.0204 | -34.7572 | -23.969
13.0265 | -34.7578 | -23.978
13.0326 | -34.7584 | -23.9875
13.0387 | -34.759 | -23.997
13.0448 | -34.7595 | -24.0062
13.0509 | -34.7601 | -24.0154
13.057 | -34.7607 | -24.0243
13.0631 | -34.7613 | -24.0334
13.0692 | -34.7619 | -24.0425
13.0753 | -34.7625 | -24.0517
13.0814 | -34.7631 | -24.0607
13.0875 | -34.7636 | -24.0695
13.0936 | -34.7642 | -24.0786
13.0997 | -34.7648 | -24.0875
13.1058 | -34.7653 | -24.0961
13.112 | -34.7658 | -24.1053
13.1181 | -34.7664 | -24.1145
13.1242 | -34.767 | -24.1238
13.1303 | -34.7676 | -24.1332
13.1364 | -34.7681 | -24.1423
13.1425 | -34.7687 | -24.1514
13.1486 | -34.7692 | -24.1604
13.1547 | -34.7698 | -24.1693
13.1608 | -34.7704 | -24.1783
13.1669 | -34.771 | -24.1874
13.173 | -34.7715 | -24.1965
13.1791 | -34.7721 | -24.2057
13.1852 | -34.7726 | -24.2151
13.1913 | -34.7732 | -24.2242
13.1974 | -34.7737 | -24.2333
13.2035 | -34.7743 | -24.2424
13.2096 | -34.7749 | -24.2516
13.2157 | -34.7755 | -24.2607
13.2218 | -34.7761 | -24.2699
13.2279 | -34.7767 | -24.279
13.234 | -34.7773 | -24.2882
13.2401 | -34.7778 | -24.2972
13.2462 | -34.7784 | -24.3063
13.2524 | -34.7789 | -24.3153
13.2585 | -34.7795 | -24.3244
13.2646 | -34.7801 | -24.3336
13.2707 | -34.7807 | -24.3428
13.2768 | -34.7812 | -24.3522
13.2829 | -34.7818 | -24.3614
13.289 | -34.7823 | -24.3705
13.2951 | -34.7829 | -24.3794
13.3012 | -34.7835 | -24.3886
13.3073 | -34.784 | -24.398
13.3134 | -34.7845 | -24.407
13.3195 | -34.7851 | -24.416
13.3256 | -34.7857 | -24.4252
13.3317 | -34.7863 | -24.4341
13.3378 | -34.7868 | -24.4433
13.3439 | -34.7874 | -24.4524
13.35 | -34.788 | -24.4615
13.3561 | -34.7885 | -24.4707
13.3622 | -34.7891 | -24.4798
13.3683 | -34.7897 | -24.4887
13.3744 | -34.7902 | -24.4977
13.3805 | -34.7907 | -24.5069
13.3866 | -34.7913 | -24.5157
13.3927 | -34.7918 | -24.5248
13.3989 | -34.7924 | -24.5336
13.405 | -34.793 | -24.5426
13.4111 | -34.7936 | -24.5516
13.4172 | -34.7941 | -24.5604
13.4233 | -34.7947 | -24.5697
13.4294 | -34.7952 | -24.5785
13.4355 | -34.7958 | -24.588
13.4416 | -34.7964 | -24.5971
13.4477 | -34.7969 | -24.6062
13.4538 | -34.7975 | -24.6152
13.4599 | -34.7981 | -24.6244
13.466 | -34.7986 | -24.6337
13.4721 | -34.7992 | -24.6428
13.4782 | -34.7998 | -24.6518
13.4843 | -34.8004 | -24.6609
13.4904 | -34.8009 | -24.6694
13.4965 | -34.8015 | -24.6788
13.5026 | -34.802 | -24.6878
13.5087 | -34.8026 | -24.6965
13.5148 | -34.8032 | -24.7057
13.5209 | -34.8037 | -24.7146
13.527 | -34.8043 | -24.7239
13.5331 | -34.8048 | -24.7328
13.5393 | -34.8054 | -24.7418
13.5454 | -34.8059 | -24.7509
13.5515 | -34.8065 | -24.76
13.5576 | -34.8071 | -24.7691
13.5637 | -34.8077 | -24.7781
13.5698 | -34.8082 | -24.7873
13.5759 | -34.8087 | -24.796
13.582 | -34.8093 | -24.8051
13.5881 | -34.8098 | -24.8141
13.5942 | -34.8104 | -24.8232
13.6003 | -34.811 | -24.8323
13.6064 | -34.8115 | -24.8415
13.6125 | -34.8121 | -24.8507
13.6186 | -34.8127 | -24.8597
13.6247 | -34.8133 | -24.8687
13.6308 | -34.8138 | -24.8774
13.6369 | -34.8144 | -24.8865
13.643 | -34.815 | -24.8954
13.6491 | -34.8155 | -24.9044
13.6552 | -34.8161 | -24.9132
13.6613 | -34.8167 | -24.9224
13.6674 | -34.8172 | -24.9315
13.6735 | -34.8178 | -24.9406
13.6796 | -34.8184 | -24.9495
13.6858 | -34.8189 | -24.9585
13.6919 | -34.8195 | -24.9676
13.698 | -34.82 | -24.9767
13.7041 | -34.8205 | -24.986
13.7102 | -34.8211 | -24.9951
13.7163 | -34.8217 | -25.0042
13.7224 | -34.8223 | -25.0131
13.7285 | -34.8229 | -25.0222
13.7346 | -34.8234 | -25.0312
13.7407 | -34.824 | -25.0403
13.7468 | -34.8245 | -25.0493
13.7529 | -34.8251 | -25.0581
13.759 | -34.8256 | -25.0673
13.7651 | -34.8262 | -25.0764
13.7712 | -34.8267 | -25.0855
13.7773 | -34.8273 | -25.0946
13.7834 | -34.8279 | -25.1036
13.7895 | -34.8285 | -25.1124
13.7956 | -34.8291 | -25.1216
13.8017 | -34.8296 | -25.1305
13.8078 | -34.8302 | -25.1395
13.8139 | -34.8308 | -25.1484
13.82 | -34.8314 | -25.1572
13.8262 | -34.832 | -25.1665
13.8323 | -34.8325 | -25.1756
13.8384 | -34.8331 | -25.1848
13.8445 | -34.8337 | -25.1939
13.8506 | -34.8342 | -25.203
13.8567 | -34.8348 | -25.2123
13.8628 | -34.8353 | -25.2214
13.8689 | -34.8359 | -25.2305
13.875 | -34.8365 | -25.2394
13.8811 | -34.8371 | -25.2483
13.8872 | -34.8377 | -25.2573
13.8933 | -34.8382 | -25.2661
13.8994 | -34.8388 | -25.275
13.9055 | -34.8394 | -25.2839
13.9116 | -34.84 | -25.2927
13.9177 | -34.8406 | -25.3014
13.9238 | -34.8412 | -25.3105
13.9299 | -34.8418 | -25.3194
13.936 | -34.8423 | -25.3283
13.9421 | -34.8428 | -25.3373
13.9482 | -34.8434 | -25.3463
13.9543 | -34.844 | -25.355
13.9604 | -34.8446 | -25.3642
13.9665 | -34.8452 | -25.3731
13.9727 | -34.8457 | -25.382
13.9788 | -34.8463 | -25.3912
13.9849 | -34.8468 | -25.4002
13.991 | -34.8473 | -25.4095
13.9971 | -34.8479 | -25.4187
14.0032 | -34.8484 | -25.4278
14.0093 | -34.849 | -25.4367
14.0154 | -34.8495 | -25.4459
14.0215 | -34.8501 | -25.4546
14.0276 | -34.8506 | -25.4634
14.0337 | -34.8513 | -25.4725
14.0398 | -34.8518 | -25.4813
14.0459 | -34.8524 | -25.4904
14.052 | -34.8529 | -25.4995
14.0581 | -34.8534 | -25.5081
14.0642 | -34.854 | -25.5173
14.0703 | -34.8546 | -25.5263
14.0764 | -34.8552 | -25.5356
14.0825 | -34.8557 | -25.5446
14.0886 | -34.8563 | -25.5532
14.0947 | -34.8568 | -25.5619
14.1008 | -34.8574 | -25.5712
14.1069 | -34.858 | -25.5801
14.1131 | -34.8585 | -25.5893
14.1192 | -34.859 | -25.598
14.1253 | -34.8596 | -25.6067
14.1314 | -34.8602 | -25.6156
14.1375 | -34.8608 | -25.6247
14.1436 | -34.8614 | -25.6336
14.1497 | -34.862 | -25.6424
14.1558 | -34.8625 | -25.6515
14.1619 | -34.8631 | -25.6604
14.168 | -34.8637 | -25.6692
14.1741 | -34.8643 | -25.678
14.1802 | -34.8649 | -25.6868
14.1863 | -34.8654 | -25.6958
14.1924 | -34.866 | -25.7048
14.1985 | -34.8666 | -25.7136
14.2046 | -34.8671 | -25.7223
14.2107 | -34.8677 | -25.7309
14.2168 | -34.8683 | -25.7401
14.2229 | -34.8689 | -25.7491
14.229 | -34.8694 | -25.758
14.2351 | -34.87 | -25.7668
14.2412 | -34.8706 | -25.7755
14.2473 | -34.8712 | -25.7844
14.2534 | -34.8718 | -25.7932
14.2596 | -34.8723 | -25.8022
14.2657 | -34.8729 | -25.8112
14.2718 | -34.8735 | -25.8203
14.2779 | -34.874 | -25.829
14.284 | -34.8746 | -25.8378
14.2901 | -34.8752 | -25.8466
14.2962 | -34.8758 | -25.8552
14.3023 | -34.8764 | -25.8642
14.3084 | -34.8769 | -25.8733
14.3145 | -34.8775 | -25.8822
14.3206 | -34.8781 | -25.8911
14.3267 | -34.8786 | -25.9
14.3328 | -34.8792 | -25.9088
14.3389 | -34.8798 | -25.9175
14.345 | -34.8803 | -25.9265
14.3511 | -34.8809 | -25.9351
14.3572 | -34.8815 | -25.9442
14.3633 | -34.882 | -25.953
14.3694 | -34.8826 | -25.9619
14.3755 | -34.8831 | -25.9705
14.3816 | -34.8837 | -25.9793
14.3877 | -34.8843 | -25.9879
14.3938 | -34.8849 | -25.997
14.4 | -34.8855 | -26.0057
14.4061 | -34.886 | -26.0146
14.4122 | -34.8865 | -26.0236
14.4183 | -34.8871 | -26.0325
14.4244 | -34.8876 | -26.0412
14.4305 | -34.8882 | -26.0503
14.4366 | -34.8888 | -26.0593
14.4427 | -34.8894 | -26.0683
14.4488 | -34.8899 | -26.0769
14.4549 | -34.8905 | -26.0856
14.461 | -34.8911 | -26.0946
14.4671 | -34.8917 | -26.1035
14.4732 | -34.8923 | -26.112
14.4793 | -34.8929 | -26.1215
14.4854 | -34.8935 | -26.13
14.4915 | -34.894 | -26.1386
14.4976 | -34.8947 | -26.1475
14.5037 | -34.8953 | -26.1565
14.5098 | -34.8958 | -26.1651
14.5159 | -34.8964 | -26.1741
14.522 | -34.897 | -26.1831
14.5281 | -34.8976 | -26.1917
14.5342 | -34.8981 | -26.2006
14.5403 | -34.8987 | -26.2095
14.5465 | -34.8993 | -26.2187
14.5526 | -34.8999 | -26.2278
14.5587 | -34.9004 | -26.2366
14.5648 | -34.901 | -26.2452
14.5709 | -34.9016 | -26.2542
14.577 | -34.9022 | -26.2631
14.5831 | -34.9027 | -26.272
14.5892 | -34.9034 | -26.2805
14.5953 | -34.9039 | -26.289
14.6014 | -34.9045 | -26.2978
14.6075 | -34.9051 | -26.3064
14.6136 | -34.9057 | -26.3151
14.6197 | -34.9063 | -26.3239
14.6258 | -34.9068 | -26.3328
14.6319 | -34.9075 | -26.3414
14.638 | -34.908 | -26.3502
14.6441 | -34.9086 | -26.3586
14.6502 | -34.9092 | -26.3672
14.6563 | -34.9097 | -26.3759
14.6624 | -34.9103 | -26.3846
14.6685 | -34.9109 | -26.3931
14.6746 | -34.9115 | -26.4019
14.6807 | -34.9121 | -26.4107
14.6869 | -34.9126 | -26.4194
14.693 | -34.9132 | -26.4281
14.6991 | -34.9138 | -26.4367
14.7052 | -34.9144 | -26.4453
14.7113 | -34.915 | -26.4543
14.7174 | -34.9155 | -26.463
14.7235 | -34.9161 | -26.4716
14.7296 | -34.9167 | -26.4803
14.7357 | -34.9173 | -26.4888
14.7418 | -34.9179 | -26.4973
14.7479 | -34.9185 | -26.5061
14.754 | -34.9191 | -26.5148
14.7601 | -34.9197 | -26.5236
14.7662 | -34.9202 | -26.5324
14.7723 | -34.9208 | -26.5409
14.7784 | -34.9214 | -26.5498
14.7845 | -34.922 | -26.5586
14.7906 | -34.9226 | -26.5673
14.7967 | -34.9231 | -26.576
14.8028 | -34.9237 | -26.5845
14.8089 | -34.9243 | -26.5931
14.815 | -34.9249 | -26.6017
14.8211 | -34.9254 | -26.6103
14.8272 | -34.926 | -26.6192
14.8334 | -34.9266 | -26.6278
14.8395 | -34.9272 | -26.6365
14.8456 | -34.9278 | -26.645
14.8517 | -34.9284 | -26.6538
14.8578 | -34.9289 | -26.6624
14.8639 | -34.9295 | -26.6713
14.87 | -34.9301 | -26.68
14.8761 | -34.9307 | -26.6884
14.8822 | -34.9313 | -26.6972
14.8883 | -34.9319 | -26.7057
14.8944 | -34.9325 | -26.7145
14.9005 | -34.9331 | -26.723
14.9066 | -34.9337 | -26.7317
14.9127 | -34.9343 | -26.7402
14.9188 | -34.9348 | -26.749
14.9249 | -34.9354 | -26.7577
14.931 | -34.936 | -26.7663
14.9371 | -34.9366 | -26.7748
14.9432 | -34.9372 | -26.7836
14.9493 | -34.9378 | -26.7922
14.9554 | -34.9384 | -26.8008
14.9615 | -34.9389 | -26.8095
14.9676 | -34.9395 | -26.8183
14.9738 | -34.9401 | -26.8268
14.9799 | -34.9407 | -26.8358
14.986 | -34.9413 | -26.8441
14.9921 | -34.9419 | -26.8526
14.9982 | -34.9425 | -26.8615
15.0043 | -34.943 | -26.8701
15.0104 | -34.9436 | -26.8788
15.0165 | -34.9443 | -26.8876
15.0226 | -34.9448 | -26.8958
15.0287 | -34.9454 | -26.9049
15.0348 | -34.946 | -26.9137
15.0409 | -34.9466 | -26.9226
15.047 | -34.9472 | -26.9314
15.0531 | -34.9479 | -26.9401
15.0592 | -34.9485 | -26.9489
15.0653 | -34.9491 | -26.9575
15.0714 | -34.9497 | -26.9659
15.0775 | -34.9503 | -26.9742
15.0836 | -34.9509 | -26.9829
15.0897 | -34.9515 | -26.9914
15.0958 | -34.9521 | -26.9999
15.1019 | -34.9527 | -27.0085
15.108 | -34.9532 | -27.0172
15.1141 | -34.9538 | -27.0258
15.1203 | -34.9545 | -27.0344
15.1264 | -34.9551 | -27.0431
15.1325 | -34.9557 | -27.0517
15.1386 | -34.9563 | -27.0603
15.1447 | -34.9569 | -27.069
15.1508 | -34.9576 | -27.0778
15.1569 | -34.9582 | -27.0867
15.163 | -34.9588 | -27.0948
15.1691 | -34.9595 | -27.1034
15.1752 | -34.96 | -27.1119
15.1813 | -34.9606 | -27.1206
15.1874 | -34.9613 | -27.1293
15.1935 | -34.9619 | -27.1375
15.1996 | -34.9626 | -27.1463
15.2057 | -34.9631 | -27.1548
15.2118 | -34.9637 | -27.1634
15.2179 | -34.9643 | -27.1718
15.224 | -34.9649 | -27.1805
15.2301 | -34.9655 | -27.1887
15.2362 | -34.9661 | -27.1973
15.2423 | -34.9667 | -27.2061
15.2484 | -34.9673 | -27.2145
15.2545 | -34.9679 | -27.2231
15.2607 | -34.9686 | -27.2317
15.2668 | -34.9692 | -27.24
15.2729 | -34.9698 | -27.2485
15.279 | -34.9705 | -27.2571
15.2851 | -34.9711 | -27.2657
15.2912 | -34.9717 | -27.2742
15.2973 | -34.9722 | -27.2827
15.3034 | -34.9729 | -27.291
15.3095 | -34.9735 | -27.2993
15.3156 | -34.9741 | -27.308
15.3217 | -34.9747 | -27.3162
15.3278 | -34.9753 | -27.3249
15.3339 | -34.976 | -27.3331
15.34 | -34.9766 | -27.3415
15.3461 | -34.9772 | -27.3501
15.3522 | -34.9777 | -27.3585
15.3583 | -34.9784 | -27.367
15.3644 | -34.979 | -27.3756
15.3705 | -34.9796 | -27.3838
15.3766 | -34.9802 | -27.3926
15.3827 | -34.9808 | -27.4009
15.3888 | -34.9814 | -27.4091
15.3949 | -34.982 | -27.4175
15.401 | -34.9826 | -27.4261
15.4072 | -34.9833 | -27.4342
15.4133 | -34.9839 | -27.4428
15.4194 | -34.9845 | -27.4514
15.4255 | -34.9851 | -27.4598
15.4316 | -34.9858 | -27.4684
15.4377 | -34.9864 | -27.4768
15.4438 | -34.987 | -27.4849
15.4499 | -34.9876 | -27.4933
15.456 | -34.9882 | -27.5018
15.4621 | -34.9888 | -27.5103
15.4682 | -34.9894 | -27.5189
15.4743 | -34.9901 | -27.5272
15.4804 | -34.9907 | -27.5358
15.4865 | -34.9913 | -27.5442
15.4926 | -34.9919 | -27.5528
15.4987 | -34.9925 | -27.5613
15.5048 | -34.9931 | -27.5697
15.5109 | -34.9938 | -27.5778
15.517 | -34.9944 | -27.5864
15.5231 | -34.995 | -27.5948
15.5292 | -34.9956 | -27.6031
15.5353 | -34.9963 | -27.6115
15.5414 | -34.9969 | -27.6201
15.5476 | -34.9975 | -27.6286
15.5537 | -34.9981 | -27.637
15.5598 | -34.9988 | -27.6452
15.5659 | -34.9994 | -27.6539
15.572 | -35 | -27.662
15.5781 | -35.0006 | -27.6704
15.5842 | -35.0013 | -27.6785
15.5903 | -35.0019 | -27.6866
15.5964 | -35.0025 | -27.6949
15.6025 | -35.0032 | -27.7031
15.6086 | -35.0038 | -27.7116
15.6147 | -35.0044 | -27.7201
15.6208 | -35.005 | -27.7282
15.6269 | -35.0057 | -27.7364
15.633 | -35.0063 | -27.7449
15.6391 | -35.0069 | -27.7534
15.6452 | -35.0075 | -27.7617
15.6513 | -35.0081 | -27.7703
15.6574 | -35.0088 | -27.7784
15.6635 | -35.0094 | -27.7868
15.6696 | -35.01 | -27.7951
15.6757 | -35.0106 | -27.8038
15.6818 | -35.0112 | -27.8124
15.688 | -35.0119 | -27.8207
15.6941 | -35.0125 | -27.8291
15.7002 | -35.0132 | -27.8374
15.7063 | -35.0138 | -27.8459
15.7124 | -35.0144 | -27.8544
15.7185 | -35.015 | -27.8627
15.7246 | -35.0156 | -27.8709
15.7307 | -35.0163 | -27.8794
15.7368 | -35.0169 | -27.888
15.7429 | -35.0175 | -27.8963
15.749 | -35.0181 | -27.9047
15.7551 | -35.0187 | -27.913
15.7612 | -35.0194 | -27.9208
15.7673 | -35.02 | -27.9291
15.7734 | -35.0206 | -27.9372
15.7795 | -35.0212 | -27.9459
15.7856 | -35.0218 | -27.9541
15.7917 | -35.0225 | -27.9622
15.7978 | -35.0231 | -27.9704
15.8039 | -35.0237 | -27.9787
15.81 | -35.0243 | -27.987
15.8161 | -35.0249 | -27.9954
15.8222 | -35.0255 | -28.0037
15.8283 | -35.0262 | -28.0118
15.8345 | -35.0268 | -28.0202
15.8406 | -35.0275 | -28.0285
15.8467 | -35.0281 | -28.0369
15.8528 | -35.0287 | -28.0452
15.8589 | -35.0294 | -28.0534
15.865 | -35.03 | -28.0615
15.8711 | -35.0306 | -28.0701
15.8772 | -35.0313 | -28.0784
15.8833 | -35.0319 | -28.0871
15.8894 | -35.0325 | -28.0954
15.8955 | -35.0332 | -28.1038
15.9016 | -35.0338 | -28.1119
15.9077 | -35.0345 | -28.1202
15.9138 | -35.0351 | -28.1285
15.9199 | -35.0357 | -28.1369
15.926 | -35.0363 | -28.1455
15.9321 | -35.037 | -28.1537
15.9382 | -35.0377 | -28.1619
15.9443 | -35.0383 | -28.1704
15.9504 | -35.039 | -28.1787
15.9565 | -35.0396 | -28.187
15.9626 | -35.0403 | -28.1953
15.9687 | -35.0409 | -28.2034
15.9749 | -35.0416 | -28.2118
15.981 | -35.0422 | -28.2201
15.9871 | -35.0428 | -28.2284
15.9932 | -35.0435 | -28.2368
15.9993 | -35.0441 | -28.245
16.0054 | -35.0447 | -28.2529
16.0115 | -35.0454 | -28.2611
16.0176 | -35.046 | -28.2695
16.0237 | -35.0467 | -28.2781
16.0298 | -35.0473 | -28.286
16.0359 | -35.0479 | -28.2942
16.042 | -35.0486 | -28.3024
16.0481 | -35.0492 | -28.3106
16.0542 | -35.0499 | -28.3191
16.0603 | -35.0505 | -28.3273
16.0664 | -35.0512 | -28.3356
16.0725 | -35.0518 | -28.3436
16.0786 | -35.0525 | -28.3521
16.0847 | -35.0531 | -28.3601
16.0908 | -35.0537 | -28.3684
16.0969 | -35.0544 | -28.3764
16.103 | -35.055 | -28.3849
16.1091 | -35.0556 | -28.3929
16.1152 | -35.0562 | -28.4013
16.1214 | -35.0569 | -28.4096
16.1275 | -35.0575 | -28.4177
16.1336 | -35.0582 | -28.4259
16.1397 | -35.0588 | -28.4342
16.1458 | -35.0595 | -28.4423
16.1519 | -35.0601 | -28.4507
16.158 | -35.0608 | -28.4591
16.1641 | -35.0614 | -28.4674
16.1702 | -35.062 | -28.4757
16.1763 | -35.0627 | -28.484
16.1824 | -35.0633 | -28.4919
16.1885 | -35.064 | -28.5001
16.1946 | -35.0646 | -28.5084
16.2007 | -35.0652 | -28.5169
16.2068 | -35.0659 | -28.525
16.2129 | -35.0665 | -28.5334
16.219 | -35.0671 | -28.5415
16.2251 | -35.0677 | -28.5499
16.2312 | -35.0683 | -28.5582
16.2373 | -35.069 | -28.5664
16.2434 | -35.0696 | -28.5746
16.2495 | -35.0703 | -28.5825
16.2556 | -35.0709 | -28.5908
16.2618 | -35.0716 | -28.5988
16.2679 | -35.0722 | -28.607
16.274 | -35.0728 | -28.6154
16.2801 | -35.0735 | -28.6234
16.2862 | -35.0741 | -28.6314
16.2923 | -35.0748 | -28.6398
16.2984 | -35.0754 | -28.6481
16.3045 | -35.0761 | -28.6563
16.3106 | -35.0767 | -28.6642
16.3167 | -35.0773 | -28.6726
16.3228 | -35.078 | -28.6808
16.3289 | -35.0786 | -28.689
16.335 | -35.0792 | -28.6973
16.3411 | -35.0798 | -28.7054
16.3472 | -35.0805 | -28.7137
16.3533 | -35.0811 | -28.7219
16.3594 | -35.0818 | -28.7299
16.3655 | -35.0824 | -28.7382
16.3716 | -35.083 | -28.7463
16.3777 | -35.0836 | -28.7544
16.3838 | -35.0843 | -28.7625
16.3899 | -35.0849 | -28.7708
16.396 | -35.0855 | -28.7791
16.4021 | -35.0862 | -28.7872
16.4083 | -35.0869 | -28.7953
16.4144 | -35.0875 | -28.8034
16.4205 | -35.0882 | -28.8117
16.4266 | -35.0888 | -28.8197
16.4327 | -35.0894 | -28.8278
16.4388 | -35.0901 | -28.836
16.4449 | -35.0907 | -28.8442
16.451 | -35.0913 | -28.8526
16.4571 | -35.0919 | -28.8609
16.4632 | -35.0926 | -28.8689
16.4693 | -35.0932 | -28.8771
16.4754 | -35.0938 | -28.8854
16.4815 | -35.0945 | -28.8937
16.4876 | -35.0951 | -28.9018
16.4937 | -35.0957 | -28.9099
16.4998 | -35.0964 | -28.918
16.5059 | -35.0971 | -28.9262
16.512 | -35.0977 | -28.9346
16.5181 | -35.0984 | -28.9429
16.5242 | -35.099 | -28.9511
16.5303 | -35.0996 | -28.959
16.5364 | -35.1002 | -28.9671
16.5425 | -35.1009 | -28.9754
16.5487 | -35.1015 | -28.9835
16.5548 | -35.1021 | -28.9916
16.5609 | -35.1027 | -29
16.567 | -35.1034 | -29.0081
16.5731 | -35.104 | -29.0163
16.5792 | -35.1046 | -29.0245
16.5853 | -35.1053 | -29.0325
16.5914 | -35.1059 | -29.0408
16.5975 | -35.1066 | -29.0491
16.6036 | -35.1072 | -29.0571
16.6097 | -35.1079 | -29.0652
16.6158 | -35.1085 | -29.0735
16.6219 | -35.1091 | -29.0816
16.628 | -35.1097 | -29.0898
16.6341 | -35.1104 | -29.0978
16.6402 | -35.111 | -29.1054
16.6463 | -35.1116 | -29.1136
16.6524 | -35.1123 | -29.1216
16.6585 | -35.1129 | -29.1299
16.6646 | -35.1135 | -29.1381
16.6707 | -35.1142 | -29.146
16.6768 | -35.1148 | -29.1541
16.6829 | -35.1155 | -29.1622
16.689 | -35.1161 | -29.1705
16.6952 | -35.1167 | -29.1787
16.7013 | -35.1173 | -29.1868
16.7074 | -35.118 | -29.1948
16.7135 | -35.1186 | -29.2029
16.7196 | -35.1192 | -29.211
16.7257 | -35.1199 | -29.219
16.7318 | -35.1205 | -29.2272
16.7379 | -35.1211 | -29.2354
16.744 | -35.1218 | -29.2432
16.7501 | -35.1224 | -29.2514
16.7562 | -35.1231 | -29.2595
16.7623 | -35.1237 | -29.2676
16.7684 | -35.1243 | -29.2759
16.7745 | -35.125 | -29.2837
16.7806 | -35.1256 | -29.2918
16.7867 | -35.1262 | -29.2999
16.7928 | -35.1269 | -29.308
16.7989 | -35.1275 | -29.3161
16.805 | -35.1281 | -29.3242
16.8111 | -35.1287 | -29.3323
16.8172 | -35.1294 | -29.3407
16.8233 | -35.13 | -29.349
16.8294 | -35.1307 | -29.3572
16.8356 | -35.1313 | -29.3655
16.8417 | -35.1319 | -29.3736
16.8478 | -35.1325 | -29.3816
16.8539 | -35.1332 | -29.3896
16.86 | -35.1338 | -29.3977
16.8661 | -35.1344 | -29.4058
16.8722 | -35.1351 | -29.4139
16.8783 | -35.1357 | -29.4218
16.8844 | -35.1364 | -29.4297
16.8905 | -35.137 | -29.438
16.8966 | -35.1376 | -29.4462
16.9027 | -35.1383 | -29.4541
16.9088 | -35.1389 | -29.4625
16.9149 | -35.1396 | -29.4709
16.921 | -35.1402 | -29.4788
16.9271 | -35.1408 | -29.4873
16.9332 | -35.1415 | -29.4954
16.9393 | -35.1421 | -29.5037
16.9454 | -35.1427 | -29.5116
16.9515 | -35.1433 | -29.5199
16.9576 | -35.1439 | -29.5278
16.9637 | -35.1446 | -29.5359
16.9698 | -35.1453 | -29.5438
16.9759 | -35.1459 | -29.5517
16.9821 | -35.1465 | -29.5599
16.9882 | -35.1471 | -29.5681
16.9943 | -35.1478 | -29.5763
17.0004 | -35.1484 | -29.5844
17.0065 | -35.149 | -29.5925
17.0126 | -35.1496 | -29.6005
17.0187 | -35.1503 | -29.6087
17.0248 | -35.1509 | -29.6166
17.0309 | -35.1515 | -29.6247
17.037 | -35.1522 | -29.6328
17.0431 | -35.1528 | -29.6407
17.0492 | -35.1534 | -29.6487
17.0553 | -35.154 | -29.6569
17.0614 | -35.1546 | -29.6649
17.0675 | -35.1552 | -29.6732
17.0736 | -35.1559 | -29.6813
17.0797 | -35.1565 | -29.689
17.0858 | -35.1571 | -29.6973
17.0919 | -35.1577 | -29.7055
17.098 | -35.1584 | -29.7133
17.1041 | -35.159 | -29.7212
17.1102 | -35.1596 | -29.7295
17.1163 | -35.1603 | -29.7376
17.1225 | -35.1609 | -29.7455
17.1286 | -35.1615 | -29.7533
17.1347 | -35.1621 | -29.7613
17.1408 | -35.1628 | -29.7694
17.1469 | -35.1634 | -29.7773
17.153 | -35.1641 | -29.7853
17.1591 | -35.1647 | -29.7931
17.1652 | -35.1653 | -29.801
17.1713 | -35.1659 | -29.809
17.1774 | -35.1665 | -29.8171
17.1835 | -35.1672 | -29.8253
17.1896 | -35.1678 | -29.8332
17.1957 | -35.1684 | -29.8413
17.2018 | -35.169 | -29.8494
17.2079 | -35.1696 | -29.8574
17.214 | -35.1702 | -29.8654
17.2201 | -35.1709 | -29.8735
17.2262 | -35.1715 | -29.8816
17.2323 | -35.1721 | -29.8896
17.2384 | -35.1727 | -29.8976
17.2445 | -35.1733 | -29.9055
17.2506 | -35.1739 | -29.9137
17.2567 | -35.1745 | -29.9218
17.2628 | -35.1752 | -29.9301
17.269 | -35.1758 | -29.9381
17.2751 | -35.1764 | -29.9461
17.2812 | -35.177 | -29.9541
17.2873 | -35.1777 | -29.9623
17.2934 | -35.1783 | -29.9703
17.2995 | -35.1789 | -29.9782
17.3056 | -35.1795 | -29.9861
17.3117 | -35.1801 | -29.9941
17.3178 | -35.1807 | -30.002
17.3239 | -35.1813 | -30.01
17.33 | -35.1819 | -30.0179
17.3361 | -35.1825 | -30.0258
17.3422 | -35.1831 | -30.0338
17.3483 | -35.1837 | -30.0419
17.3544 | -35.1843 | -30.0498
17.3605 | -35.185 | -30.0578
17.3666 | -35.1856 | -30.066
17.3727 | -35.1862 | -30.0739
17.3788 | -35.1868 | -30.0817
17.3849 | -35.1875 | -30.0897
17.391 | -35.1881 | -30.0975
17.3971 | -35.1887 | -30.1056
17.4032 | -35.1893 | -30.1139
17.4094 | -35.1899 | -30.1217
17.4155 | -35.1905 | -30.1297
17.4216 | -35.1911 | -30.1377
17.4277 | -35.1918 | -30.1459
17.4338 | -35.1924 | -30.1539
17.4399 | -35.193 | -30.1619
17.446 | -35.1936 | -30.1697
17.4521 | -35.1942 | -30.1775
17.4582 | -35.1948 | -30.1854
17.4643 | -35.1954 | -30.1935
17.4704 | -35.196 | -30.2014
17.4765 | -35.1967 | -30.2094
17.4826 | -35.1973 | -30.2173
17.4887 | -35.1979 | -30.2255
17.4948 | -35.1985 | -30.2334
17.5009 | -35.1991 | -30.2412
17.507 | -35.1997 | -30.2493
17.5131 | -35.2004 | -30.2573
17.5192 | -35.2009 | -30.2651
17.5253 | -35.2015 | -30.2732
17.5314 | -35.2021 | -30.2812
17.5375 | -35.2027 | -30.2892
17.5436 | -35.2033 | -30.2973
17.5497 | -35.2039 | -30.3055
17.5559 | -35.2045 | -30.3132
17.562 | -35.2051 | -30.321
17.5681 | -35.2058 | -30.3289
17.5742 | -35.2064 | -30.337
17.5803 | -35.207 | -30.3447
17.5864 | -35.2076 | -30.3524
17.5925 | -35.2083 | -30.3607
17.5986 | -35.2089 | -30.3684
17.6047 | -35.2095 | -30.3766
17.6108 | -35.2101 | -30.3843
17.6169 | -35.2108 | -30.3922
17.623 | -35.2113 | -30.3999
17.6291 | -35.212 | -30.4077
17.6352 | -35.2126 | -30.4155
17.6413 | -35.2132 | -30.4236
17.6474 | -35.2138 | -30.4316
17.6535 | -35.2144 | -30.4396
17.6596 | -35.215 | -30.4475
17.6657 | -35.2156 | -30.4553
17.6718 | -35.2162 | -30.4634
17.6779 | -35.2168 | -30.4712
17.684 | -35.2174 | -30.4792
17.6901 | -35.2181 | -30.487
17.6963 | -35.2187 | -30.4947
17.7024 | -35.2193 | -30.5025
17.7085 | -35.2199 | -30.5103
17.7146 | -35.2205 | -30.5182
17.7207 | -35.2211 | -30.5259
17.7268 | -35.2217 | -30.5337
17.7329 | -35.2223 | -30.5419
17.739 | -35.2229 | -30.5499
17.7451 | -35.2236 | -30.5578
17.7512 | -35.2242 | -30.5657
17.7573 | -35.2248 | -30.5736
17.7634 | -35.2254 | -30.5812
17.7695 | -35.226 | -30.5892
17.7756 | -35.2266 | -30.597
17.7817 | -35.2272 | -30.6047
17.7878 | -35.2278 | -30.6127
17.7939 | -35.2284 | -30.6204
17.8 | -35.2291 | -30.6282
17.8061 | -35.2296 | -30.6362
17.8122 | -35.2303 | -30.644
17.8183 | -35.2309 | -30.652
17.8244 | -35.2315 | -30.6598
17.8305 | -35.2321 | -30.6677
17.8366 | -35.2327 | -30.6752
17.8428 | -35.2333 | -30.6833
17.8489 | -35.2339 | -30.691
17.855 | -35.2345 | -30.6986
17.8611 | -35.2352 | -30.7063
17.8672 | -35.2358 | -30.714
17.8733 | -35.2364 | -30.7221
17.8794 | -35.237 | -30.73
17.8855 | -35.2376 | -30.7378
17.8916 | -35.2382 | -30.7455
17.8977 | -35.2388 | -30.7533
17.9038 | -35.2394 | -30.7609
17.9099 | -35.24 | -30.7687
17.916 | -35.2406 | -30.7766
17.9221 | -35.2412 | -30.7844
17.9282 | -35.2418 | -30.7923
17.9343 | -35.2425 | -30.8002
17.9404 | -35.243 | -30.8078
17.9465 | -35.2437 | -30.8158
17.9526 | -35.2442 | -30.8238
17.9587 | -35.2449 | -30.8316
17.9648 | -35.2455 | -30.8394
17.9709 | -35.2461 | -30.8472
17.977 | -35.2467 | -30.8548
17.9832 | -35.2473 | -30.8627
17.9893 | -35.2479 | -30.8707
17.9954 | -35.2485 | -30.8787
18.0015 | -35.2491 | -30.8865
18.0076 | -35.2497 | -30.8939
18.0137 | -35.2503 | -30.902
18.0198 | -35.2509 | -30.9098
18.0259 | -35.2514 | -30.9177
18.032 | -35.252 | -30.9253
18.0381 | -35.2526 | -30.9329
18.0442 | -35.2532 | -30.9408
18.0503 | -35.2538 | -30.9487
18.0564 | -35.2544 | -30.9565
18.0625 | -35.255 | -30.9645
18.0686 | -35.2556 | -30.9722
18.0747 | -35.2562 | -30.98
18.0808 | -35.2568 | -30.9875
18.0869 | -35.2575 | -30.9954
18.093 | -35.2581 | -31.0032
18.0991 | -35.2587 | -31.0112
18.1052 | -35.2593 | -31.0192
18.1113 | -35.2599 | -31.0268
18.1174 | -35.2605 | -31.0343
18.1236 | -35.2611 | -31.0422
18.1297 | -35.2617 | -31.0499
18.1358 | -35.2624 | -31.0577
18.1419 | -35.263 | -31.0656
18.148 | -35.2636 | -31.0735
18.1541 | -35.2643 | -31.081
18.1602 | -35.2649 | -31.0889
18.1663 | -35.2655 | -31.0966
18.1724 | -35.2661 | -31.1046
18.1785 | -35.2667 | -31.1125
18.1846 | -35.2673 | -31.1201
18.1907 | -35.2679 | -31.1282
18.1968 | -35.2684 | -31.1361
18.2029 | -35.2691 | -31.1435
18.209 | -35.2697 | -31.1516
18.2151 | -35.2703 | -31.1595
18.2212 | -35.2709 | -31.167
18.2273 | -35.2716 | -31.1748
18.2334 | -35.2722 | -31.1826
18.2395 | -35.2727 | -31.1902
18.2456 | -35.2734 | -31.1979
18.2517 | -35.2739 | -31.2057
18.2578 | -35.2745 | -31.2131
18.2639 | -35.2751 | -31.2212
18.2701 | -35.2758 | -31.2289
18.2762 | -35.2764 | -31.2368
18.2823 | -35.277 | -31.2446
18.2884 | -35.2776 | -31.2525
18.2945 | -35.2782 | -31.26
18.3006 | -35.2788 | -31.2678
18.3067 | -35.2794 | -31.2754
18.3128 | -35.2801 | -31.2832
18.3189 | -35.2807 | -31.2911
18.325 | -35.2813 | -31.2984
18.3311 | -35.2819 | -31.3063
18.3372 | -35.2825 | -31.3138
18.3433 | -35.283 | -31.3214
18.3494 | -35.2836 | -31.3291
18.3555 | -35.2842 | -31.3366
18.3616 | -35.2849 | -31.3441
18.3677 | -35.2855 | -31.3516
18.3738 | -35.2861 | -31.3592
18.3799 | -35.2867 | -31.3668
18.386 | -35.2872 | -31.3744
18.3921 | -35.2879 | -31.3821
18.3982 | -35.2884 | -31.3896
18.4043 | -35.289 | -31.3972
18.4105 | -35.2896 | -31.4049
18.4166 | -35.2902 | -31.4125
18.4227 | -35.2908 | -31.4203
18.4288 | -35.2914 | -31.4282
18.4349 | -35.292 | -31.4357
18.441 | -35.2927 | -31.4434
18.4471 | -35.2933 | -31.4512
18.4532 | -35.2939 | -31.459
18.4593 | -35.2945 | -31.4666
18.4654 | -35.295 | -31.4739
18.4715 | -35.2956 | -31.4817
18.4776 | -35.2963 | -31.4895
18.4837 | -35.2969 | -31.4973
18.4898 | -35.2975 | -31.5051
18.4959 | -35.2981 | -31.5128
18.502 | -35.2987 | -31.5206
18.5081 | -35.2994 | -31.5285
18.5142 | -35.2999 | -31.5363
18.5203 | -35.3005 | -31.5437
18.5264 | -35.3011 | -31.5511
18.5325 | -35.3018 | -31.559
18.5386 | -35.3024 | -31.5668
18.5447 | -35.303 | -31.5743
18.5508 | -35.3036 | -31.5819
18.557 | -35.3042 | -31.5897
18.5631 | -35.3048 | -31.5974
18.5692 | -35.3054 | -31.605
18.5753 | -35.306 | -31.6124
18.5814 | -35.3066 | -31.6202
18.5875 | -35.3072 | -31.6281
18.5936 | -35.3078 | -31.6361
18.5997 | -35.3084 | -31.6438
18.6058 | -35.309 | -31.6513
18.6119 | -35.3096 | -31.6591
18.618 | -35.3102 | -31.6669
18.6241 | -35.3108 | -31.6747
18.6302 | -35.3115 | -31.6826
18.6363 | -35.3121 | -31.6901
18.6424 | -35.3127 | -31.6975
18.6485 | -35.3132 | -31.7052
18.6546 | -35.3139 | -31.7127
18.6607 | -35.3144 | -31.7203
18.6668 | -35.3151 | -31.7279
18.6729 | -35.3157 | -31.7357
18.679 | -35.3163 | -31.743
18.6851 | -35.3169 | -31.7505
18.6912 | -35.3175 | -31.7582
18.6974 | -35.3181 | -31.7658
18.7035 | -35.3187 | -31.7734
18.7096 | -35.3193 | -31.7811
18.7157 | -35.3199 | -31.7886
18.7218 | -35.3205 | -31.7963
18.7279 | -35.3212 | -31.8037
18.734 | -35.3218 | -31.8113
18.7401 | -35.3224 | -31.8187
18.7462 | -35.323 | -31.8258
18.7523 | -35.3236 | -31.8333
18.7584 | -35.3243 | -31.8411
18.7645 | -35.3249 | -31.8487
18.7706 | -35.3255 | -31.8563
18.7767 | -35.3261 | -31.8638
18.7828 | -35.3267 | -31.871
18.7889 | -35.3273 | -31.8785
18.795 | -35.3279 | -31.886
18.8011 | -35.3285 | -31.8936
18.8072 | -35.3291 | -31.9012
18.8133 | -35.3297 | -31.9089
18.8194 | -35.3303 | -31.9163
18.8255 | -35.3309 | -31.9238
18.8316 | -35.3315 | -31.9316
18.8377 | -35.3321 | -31.9391
18.8439 | -35.3327 | -31.9466
18.85 | -35.3333 | -31.9543
18.8561 | -35.3339 | -31.9617
18.8622 | -35.3345 | -31.9694
18.8683 | -35.3351 | -31.9769
18.8744 | -35.3357 | -31.9843
18.8805 | -35.3363 | -31.992
18.8866 | -35.3369 | -31.9998
18.8927 | -35.3376 | -32.0072
18.8988 | -35.3382 | -32.0147
18.9049 | -35.3388 | -32.0222
18.911 | -35.3394 | -32.0298
18.9171 | -35.34 | -32.0373
18.9232 | -35.3406 | -32.0446
18.9293 | -35.3412 | -32.0522
18.9354 | -35.3418 | -32.0595
18.9415 | -35.3424 | -32.0669
18.9476 | -35.343 | -32.0745
18.9537 | -35.3436 | -32.0823
18.9598 | -35.3443 | -32.0893
18.9659 | -35.3449 | -32.0968
18.972 | -35.3455 | -32.1042
18.9781 | -35.3461 | -32.1117
18.9843 | -35.3467 | -32.1193
18.9904 | -35.3473 | -32.1269
18.9965 | -35.3479 | -32.1342
19.0026 | -35.3485 | -32.1418
19.0087 | -35.3491 | -32.1494
19.0148 | -35.3497 | -32.1568
19.0209 | -35.3503 | -32.1646
19.027 | -35.351 | -32.1723
19.0331 | -35.3516 | -32.1797
19.0392 | -35.3522 | -32.1873
19.0453 | -35.3528 | -32.1948
19.0514 | -35.3534 | -32.2023
19.0575 | -35.354 | -32.2099
19.0636 | -35.3546 | -32.2168
19.0697 | -35.3553 | -32.2246
19.0758 | -35.3559 | -32.232
19.0819 | -35.3565 | -32.2397
19.088 | -35.3571 | -32.247
19.0941 | -35.3577 | -32.2542
19.1002 | -35.3583 | -32.2614
19.1063 | -35.3589 | -32.269
19.1124 | -35.3596 | -32.2765
19.1185 | -35.3601 | -32.2841
19.1246 | -35.3607 | -32.2915
19.1308 | -35.3613 | -32.2987
19.1369 | -35.362 | -32.306
19.143 | -35.3625 | -32.3135
19.1491 | -35.3632 | -32.321
19.1552 | -35.3638 | -32.3288
19.1613 | -35.3644 | -32.3363
19.1674 | -35.365 | -32.3438
19.1735 | -35.3656 | -32.3511
19.1796 | -35.3662 | -32.3584
19.1857 | -35.3668 | -32.366
19.1918 | -35.3674 | -32.3736
19.1979 | -35.368 | -32.3812
19.204 | -35.3686 | -32.3883
19.2101 | -35.3692 | -32.3958
19.2162 | -35.3698 | -32.4035
19.2223 | -35.3704 | -32.4106
19.2284 | -35.371 | -32.4183
19.2345 | -35.3716 | -32.4259
19.2406 | -35.3723 | -32.4333
19.2467 | -35.3728 | -32.4409
19.2528 | -35.3734 | -32.4483
19.2589 | -35.374 | -32.4557
19.265 | -35.3746 | -32.4632
19.2712 | -35.3752 | -32.4706
19.2773 | -35.3758 | -32.4778
19.2834 | -35.3764 | -32.4851
19.2895 | -35.3771 | -32.4926
19.2956 | -35.3777 | -32.5002
19.3017 | -35.3783 | -32.5075
19.3078 | -35.3789 | -32.5149
19.3139 | -35.3795 | -32.5222
19.32 | -35.3802 | -32.5295
19.3261 | -35.3807 | -32.5369
19.3322 | -35.3814 | -32.5443
19.3383 | -35.382 | -32.5516
19.3444 | -35.3826 | -32.5587
19.3505 | -35.3832 | -32.5661
19.3566 | -35.3838 | -32.5735
19.3627 | -35.3844 | -32.5811
19.3688 | -35.385 | -32.5887
19.3749 | -35.3856 | -32.5963
19.381 | -35.3862 | -32.6032
19.3871 | -35.3868 | -32.6106
19.3932 | -35.3874 | -32.6182
19.3993 | -35.388 | -32.6256
19.4054 | -35.3887 | -32.6332
19.4115 | -35.3893 | -32.6406
19.4177 | -35.3899 | -32.6478
19.4238 | -35.3905 | -32.6551
19.4299 | -35.3911 | -32.6624
19.436 | -35.3917 | -32.6701
19.4421 | -35.3923 | -32.6771
19.4482 | -35.393 | -32.6846
19.4543 | -35.3936 | -32.6915
19.4604 | -35.3942 | -32.6987
19.4665 | -35.3948 | -32.7061
19.4726 | -35.3954 | -32.7133
19.4787 | -35.396 | -32.7206
19.4848 | -35.3966 | -32.728
19.4909 | -35.3972 | -32.7352
19.497 | -35.3978 | -32.7425
19.5031 | -35.3984 | -32.7498
19.5092 | -35.399 | -32.7571
19.5153 | -35.3996 | -32.7646
19.5214 | -35.4003 | -32.7716
19.5275 | -35.4008 | -32.7789
19.5336 | -35.4014 | -32.7862
19.5397 | -35.4021 | -32.7936
19.5458 | -35.4027 | -32.8009
19.5519 | -35.4033 | -32.8081
19.5581 | -35.4039 | -32.8151
19.5642 | -35.4045 | -32.8224
19.5703 | -35.4051 | -32.8297
19.5764 | -35.4057 | -32.837
19.5825 | -35.4063 | -32.8443
19.5886 | -35.4069 | -32.8516
19.5947 | -35.4075 | -32.8587
19.6008 | -35.4081 | -32.8661
19.6069 | -35.4087 | -32.8733
19.613 | -35.4093 | -32.8808
19.6191 | -35.4099 | -32.8881
19.6252 | -35.4105 | -32.8954
19.6313 | -35.4111 | -32.9024
19.6374 | -35.4117 | -32.9098
19.6435 | -35.4123 | -32.9169
19.6496 | -35.4129 | -32.9241
19.6557 | -35.4135 | -32.9314
19.6618 | -35.4142 | -32.9384
19.6679 | -35.4148 | -32.9457
19.674 | -35.4154 | -32.9529
19.6801 | -35.416 | -32.9601
19.6862 | -35.4166 | -32.9674
19.6923 | -35.4172 | -32.9745
19.6984 | -35.4178 | -32.9815
19.7046 | -35.4184 | -32.9888
19.7107 | -35.4191 | -32.996
19.7168 | -35.4196 | -33.0035
19.7229 | -35.4202 | -33.0109
19.729 | -35.4208 | -33.0182
19.7351 | -35.4214 | -33.0251
19.7412 | -35.422 | -33.0324
19.7473 | -35.4226 | -33.0397
19.7534 | -35.4232 | -33.047
19.7595 | -35.4238 | -33.0544
19.7656 | -35.4245 | -33.0619
19.7717 | -35.4251 | -33.0688
19.7778 | -35.4256 | -33.0763
19.7839 | -35.4262 | -33.0833
19.79 | -35.4269 | -33.0906
19.7961 | -35.4274 | -33.098
19.8022 | -35.428 | -33.1052
19.8083 | -35.4287 | -33.1125
19.8144 | -35.4293 | -33.1197
19.8205 | -35.4298 | -33.127
19.8266 | -35.4304 | -33.1342
19.8327 | -35.431 | -33.1416
19.8388 | -35.4316 | -33.1489
19.845 | -35.4323 | -33.156
19.8511 | -35.4329 | -33.1634
19.8572 | -35.4334 | -33.1707
19.8633 | -35.4341 | -33.1779
19.8694 | -35.4347 | -33.1853
19.8755 | -35.4353 | -33.1924
19.8816 | -35.4359 | -33.1997
19.8877 | -35.4365 | -33.2069
19.8938 | -35.4371 | -33.2141
19.8999 | -35.4377 | -33.2213
19.906 | -35.4384 | -33.2286
19.9121 | -35.439 | -33.2356
19.9182 | -35.4396 | -33.2426
19.9243 | -35.4402 | -33.25
19.9304 | -35.4408 | -33.2575
19.9365 | -35.4414 | -33.2649
19.9426 | -35.442 | -33.2718
19.9487 | -35.4426 | -33.2788
19.9548 | -35.4432 | -33.2863
19.9609 | -35.4439 | -33.2934
19.967 | -35.4445 | -33.3005
19.9731 | -35.445 | -33.3077
19.9792 | -35.4456 | -33.3145
19.9853 | -35.4463 | -33.3217
19.9915 | -35.4469 | -33.3291
19.9976 | -35.4475 | -33.3362
20.0037 | -35.4481 | -33.3433
20.0098 | -35.4487 | -33.3506
20.0159 | -35.4493 | -33.3576
20.022 | -35.4499 | -33.3651
20.0281 | -35.4506 | -33.3725
20.0342 | -35.4511 | -33.3795
20.0403 | -35.4518 | -33.3864
20.0464 | -35.4524 | -33.3939
20.0525 | -35.453 | -33.4009
20.0586 | -35.4536 | -33.4079
20.0647 | -35.4542 | -33.4154
20.0708 | -35.4549 | -33.4229
20.0769 | -35.4554 | -33.43
20.083 | -35.456 | -33.4369
20.0891 | -35.4566 | -33.4441
20.0952 | -35.4572 | -33.4511
20.1013 | -35.4578 | -33.4583
20.1074 | -35.4584 | -33.4657
20.1135 | -35.459 | -33.4728
20.1196 | -35.4596 | -33.4797
20.1257 | -35.4602 | -33.487
20.1319 | -35.4608 | -33.4942
20.138 | -35.4614 | -33.5012
20.1441 | -35.462 | -33.5084
20.1502 | -35.4626 | -33.5158
20.1563 | -35.4632 | -33.5225
20.1624 | -35.4638 | -33.5296
20.1685 | -35.4644 | -33.5369
20.1746 | -35.465 | -33.5442
20.1807 | -35.4656 | -33.5514
20.1868 | -35.4662 | -33.5585
20.1929 | -35.4668 | -33.5654
20.199 | -35.4674 | -33.5727
20.2051 | -35.468 | -33.5797
20.2112 | -35.4686 | -33.5867
20.2173 | -35.4692 | -33.5939
20.2234 | -35.4698 | -33.6011
20.2295 | -35.4704 | -33.608
20.2356 | -35.471 | -33.6154
20.2417 | -35.4716 | -33.6224
20.2478 | -35.4722 | -33.6295
20.2539 | -35.4727 | -33.6368
20.26 | -35.4733 | -33.6437
20.2661 | -35.4739 | -33.651
20.2723 | -35.4745 | -33.6581
20.2784 | -35.4751 | -33.6652
20.2845 | -35.4757 | -33.6723
20.2906 | -35.4764 | -33.6795
20.2967 | -35.477 | -33.6865
20.3028 | -35.4776 | -33.6935
20.3089 | -35.4782 | -33.7007
20.315 | -35.4788 | -33.7077
20.3211 | -35.4793 | -33.7149
20.3272 | -35.48 | -33.7219
20.3333 | -35.4806 | -33.7286
20.3394 | -35.4812 | -33.7356
20.3455 | -35.4818 | -33.7426
20.3516 | -35.4824 | -33.7496
20.3577 | -35.483 | -33.7566
20.3638 | -35.4836 | -33.7637
20.3699 | -35.4842 | -33.7705
20.376 | -35.4848 | -33.7776
20.3821 | -35.4854 | -33.7846
20.3882 | -35.486 | -33.7917
20.3943 | -35.4866 | -33.799
20.4004 | -35.4872 | -33.8057
20.4065 | -35.4878 | -33.8126
20.4126 | -35.4884 | -33.8197
20.4188 | -35.489 | -33.827
20.4249 | -35.4896 | -33.8341
20.431 | -35.4902 | -33.8412
20.4371 | -35.4908 | -33.8478
20.4432 | -35.4914 | -33.8549
20.4493 | -35.492 | -33.8619
20.4554 | -35.4926 | -33.8689
20.4615 | -35.4932 | -33.8761
20.4676 | -35.4938 | -33.8832
20.4737 | -35.4944 | -33.89
20.4798 | -35.495 | -33.8971
20.4859 | -35.4956 | -33.9041
20.492 | -35.4962 | -33.9112
20.4981 | -35.4968 | -33.9181
20.5042 | -35.4974 | -33.925
20.5103 | -35.498 | -33.9318
20.5164 | -35.4986 | -33.9388
20.5225 | -35.4992 | -33.946
20.5286 | -35.4997 | -33.9533
20.5347 | -35.5003 | -33.9602
20.5408 | -35.5009 | -33.9671
20.5469 | -35.5015 | -33.9741
20.553 | -35.5022 | -33.9811
20.5592 | -35.5027 | -33.9881
20.5653 | -35.5033 | -33.9952
20.5714 | -35.5039 | -34.0019
20.5775 | -35.5045 | -34.0087
20.5836 | -35.5051 | -34.0159
20.5897 | -35.5057 | -34.0229
20.5958 | -35.5063 | -34.0299
20.6019 | -35.5069 | -34.0368
20.608 | -35.5075 | -34.0439
20.6141 | -35.5081 | -34.0507
20.6202 | -35.5087 | -34.0578
20.6263 | -35.5092 | -34.0646
20.6324 | -35.5098 | -34.0715
20.6385 | -35.5104 | -34.0784
20.6446 | -35.511 | -34.0855
20.6507 | -35.5116 | -34.0923
20.6568 | -35.5122 | -34.0992
20.6629 | -35.5128 | -34.1061
20.669 | -35.5134 | -34.113
20.6751 | -35.514 | -34.1199
20.6812 | -35.5146 | -34.1265
20.6873 | -35.5152 | -34.1336
20.6934 | -35.5158 | -34.1407
20.6995 | -35.5163 | -34.1477
20.7057 | -35.5169 | -34.1546
20.7118 | -35.5175 | -34.1613
20.7179 | -35.5181 | -34.1679
20.724 | -35.5187 | -34.1747
20.7301 | -35.5193 | -34.1818
20.7362 | -35.5199 | -34.1889
20.7423 | -35.5205 | -34.1956
20.7484 | -35.5211 | -34.2026
20.7545 | -35.5217 | -34.2092
20.7606 | -35.5223 | -34.2163
20.7667 | -35.5229 | -34.2232
20.7728 | -35.5236 | -34.2302
20.7789 | -35.5242 | -34.2371
20.785 | -35.5248 | -34.2442
20.7911 | -35.5253 | -34.2509
20.7972 | -35.5259 | -34.2578
20.8033 | -35.5265 | -34.2646
20.8094 | -35.5271 | -34.2716
20.8155 | -35.5277 | -34.2786
20.8216 | -35.5283 | -34.2857
20.8277 | -35.5289 | -34.2922
20.8338 | -35.5294 | -34.2991
20.8399 | -35.5301 | -34.3058
20.8461 | -35.5307 | -34.3125
20.8522 | -35.5313 | -34.3194
20.8583 | -35.5319 | -34.3261
20.8644 | -35.5325 | -34.3329
20.8705 | -35.5331 | -34.34
20.8766 | -35.5337 | -34.3472
20.8827 | -35.5343 | -34.3541
20.8888 | -35.5348 | -34.3609
20.8949 | -35.5354 | -34.3676
20.901 | -35.536 | -34.3745
20.9071 | -35.5366 | -34.3813
20.9132 | -35.5372 | -34.3882
20.9193 | -35.5378 | -34.3952
20.9254 | -35.5384 | -34.4021
20.9315 | -35.539 | -34.4086
20.9376 | -35.5395 | -34.4155
20.9437 | -35.5401 | -34.4224
20.9498 | -35.5407 | -34.4292
20.9559 | -35.5414 | -34.4359
20.962 | -35.542 | -34.4429
20.9681 | -35.5425 | -34.4496
20.9742 | -35.5431 | -34.4564
20.9803 | -35.5437 | -34.4633
20.9864 | -35.5443 | -34.4702
20.9926 | -35.5449 | -34.4771
20.9987 | -35.5455 | -34.4836
21.0048 | -35.5461 | -34.4906
21.0109 | -35.5466 | -34.4973
21.017 | -35.5472 | -34.5041
21.0231 | -35.5478 | -34.5109
21.0292 | -35.5484 | -34.5178
21.0353 | -35.549 | -34.5243
21.0414 | -35.5496 | -34.5311
21.0475 | -35.5502 | -34.538
21.0536 | -35.5508 | -34.5448
21.0597 | -35.5514 | -34.5515
21.0658 | -35.5519 | -34.5581
21.0719 | -35.5525 | -34.5649
21.078 | -35.5531 | -34.5719
21.0841 | -35.5537 | -34.5785
21.0902 | -35.5542 | -34.5854
21.0963 | -35.5547 | -34.5922
21.1024 | -35.5553 | -34.599
21.1085 | -35.5559 | -34.6053
21.1146 | -35.5565 | -34.6122
21.1207 | -35.5571 | -34.6189
21.1268 | -35.5577 | -34.6257
21.133 | -35.5583 | -34.6324
21.1391 | -35.5588 | -34.6388
21.1452 | -35.5594 | -34.6457
21.1513 | -35.56 | -34.6524
21.1574 | -35.5606 | -34.6595
21.1635 | -35.5612 | -34.666
21.1696 | -35.5618 | -34.6727
21.1757 | -35.5624 | -34.6791
21.1818 | -35.563 | -34.686
21.1879 | -35.5635 | -34.6928
21.194 | -35.5641 | -34.6995
21.2001 | -35.5647 | -34.7062
21.2062 | -35.5653 | -34.7129
21.2123 | -35.5659 | -34.7195
21.2184 | -35.5664 | -34.7262
21.2245 | -35.567 | -34.733
21.2306 | -35.5676 | -34.7396
21.2367 | -35.5682 | -34.7464
21.2428 | -35.5688 | -34.753
21.2489 | -35.5694 | -34.7594
21.255 | -35.57 | -34.7663
21.2611 | -35.5706 | -34.7731
21.2672 | -35.5712 | -34.7799
21.2733 | -35.5717 | -34.7867
21.2795 | -35.5723 | -34.7932
21.2856 | -35.5729 | -34.8
21.2917 | -35.5735 | -34.8069
21.2978 | -35.5741 | -34.8137
21.3039 | -35.5747 | -34.8204
21.31 | -35.5752 | -34.827
21.3161 | -35.5758 | -34.8335
21.3222 | -35.5764 | -34.8401
21.3283 | -35.577 | -34.8469
21.3344 | -35.5776 | -34.8535
21.3405 | -35.5781 | -34.8604
21.3466 | -35.5788 | -34.8671
21.3527 | -35.5793 | -34.8735
21.3588 | -35.5799 | -34.8802
21.3649 | -35.5805 | -34.8867
21.371 | -35.5811 | -34.8933
21.3771 | -35.5817 | -34.8999
21.3832 | -35.5822 | -34.9064
21.3893 | -35.5828 | -34.9127
21.3954 | -35.5834 | -34.9193
21.4015 | -35.5839 | -34.9259
21.4076 | -35.5845 | -34.9327
21.4137 | -35.5851 | -34.9394
21.4199 | -35.5857 | -34.9462
21.426 | -35.5863 | -34.9529
21.4321 | -35.5869 | -34.9596
21.4382 | -35.5874 | -34.9662
21.4443 | -35.588 | -34.9729
21.4504 | -35.5886 | -34.9796
21.4565 | -35.5891 | -34.986
21.4626 | -35.5897 | -34.9928
21.4687 | -35.5903 | -34.9995
21.4748 | -35.5909 | -35.0061
21.4809 | -35.5915 | -35.0127
21.487 | -35.5921 | -35.0193
21.4931 | -35.5927 | -35.0256
21.4992 | -35.5933 | -35.0321
21.5053 | -35.5938 | -35.039
21.5114 | -35.5944 | -35.0454
21.5175 | -35.5949 | -35.0521
21.5236 | -35.5955 | -35.0587
21.5297 | -35.5961 | -35.065
21.5358 | -35.5966 | -35.0714
21.5419 | -35.5972 | -35.0784
21.548 | -35.5977 | -35.0851
21.5541 | -35.5983 | -35.0918
21.5602 | -35.5989 | -35.0985
21.5664 | -35.5995 | -35.1049
21.5725 | -35.6 | -35.1115
21.5786 | -35.6005 | -35.1181
21.5847 | -35.6011 | -35.1248
21.5908 | -35.6016 | -35.1314
21.5969 | -35.6022 | -35.1377
21.603 | -35.6028 | -35.1444
21.6091 | -35.6034 | -35.1512
21.6152 | -35.604 | -35.1579
21.6213 | -35.6046 | -35.1647
21.6274 | -35.6052 | -35.1713
21.6335 | -35.6057 | -35.1777
21.6396 | -35.6063 | -35.1845
21.6457 | -35.6069 | -35.1912
21.6518 | -35.6074 | -35.1978
21.6579 | -35.608 | -35.2041
21.664 | -35.6086 | -35.211
21.6701 | -35.6091 | -35.2174
21.6762 | -35.6097 | -35.224
21.6823 | -35.6103 | -35.2307
21.6884 | -35.6108 | -35.2373
21.6945 | -35.6114 | -35.2439
21.7006 | -35.6119 | -35.2506
21.7068 | -35.6125 | -35.2569
21.7129 | -35.6131 | -35.2635
21.719 | -35.6137 | -35.2702
21.7251 | -35.6142 | -35.2767
21.7312 | -35.6148 | -35.2831
21.7373 | -35.6154 | -35.2895
21.7434 | -35.6159 | -35.2961
21.7495 | -35.6165 | -35.3026
21.7556 | -35.6171 | -35.3094
21.7617 | -35.6177 | -35.316
21.7678 | -35.6182 | -35.3227
21.7739 | -35.6188 | -35.3291
21.78 | -35.6194 | -35.3356
21.7861 | -35.62 | -35.3422
21.7922 | -35.6205 | -35.3488
21.7983 | -35.6211 | -35.3554
21.8044 | -35.6217 | -35.3621
21.8105 | -35.6223 | -35.3684
21.8166 | -35.6229 | -35.375
21.8227 | -35.6234 | -35.3815
21.8288 | -35.624 | -35.388
21.8349 | -35.6245 | -35.3943
21.841 | -35.6251 | -35.401
21.8471 | -35.6257 | -35.4074
21.8533 | -35.6263 | -35.4138
21.8594 | -35.6268 | -35.4207
21.8655 | -35.6274 | -35.4274
21.8716 | -35.628 | -35.4341
21.8777 | -35.6286 | -35.4404
21.8838 | -35.6292 | -35.4468
21.8899 | -35.6297 | -35.4533
21.896 | -35.6303 | -35.46
21.9021 | -35.6309 | -35.4667
21.9082 | -35.6314 | -35.4731
21.9143 | -35.632 | -35.4793
21.9204 | -35.6325 | -35.4859
21.9265 | -35.6331 | -35.4923
21.9326 | -35.6337 | -35.499
21.9387 | -35.6342 | -35.5056
21.9448 | -35.6348 | -35.5122
21.9509 | -35.6354 | -35.5185
21.957 | -35.636 | -35.525
21.9631 | -35.6365 | -35.5318
21.9692 | -35.6371 | -35.5384
21.9753 | -35.6376 | -35.5448
21.9814 | -35.6382 | -35.5513
21.9875 | -35.6387 | -35.5576
21.9937 | -35.6393 | -35.5642
21.9998 | -35.6399 | -35.5709
22.0059 | -35.6405 | -35.5776
22.012 | -35.641 | -35.5841
22.0181 | -35.6416 | -35.5904
22.0242 | -35.6422 | -35.5969
22.0303 | -35.6427 | -35.6036
22.0364 | -35.6433 | -35.6101
22.0425 | -35.6439 | -35.6166
22.0486 | -35.6444 | -35.6231
22.0547 | -35.6449 | -35.6294
22.0608 | -35.6455 | -35.6358
22.0669 | -35.646 | -35.6421
22.073 | -35.6466 | -35.6485
22.0791 | -35.6472 | -35.6551
22.0852 | -35.6478 | -35.6616
22.0913 | -35.6483 | -35.6681
22.0974 | -35.6488 | -35.6744
22.1035 | -35.6494 | -35.6807
22.1096 | -35.6499 | -35.6874
22.1157 | -35.6505 | -35.6939
22.1218 | -35.6511 | -35.7006
22.1279 | -35.6516 | -35.7068
22.134 | -35.6522 | -35.7131
22.1402 | -35.6528 | -35.7195
22.1463 | -35.6533 | -35.726
22.1524 | -35.6539 | -35.7326
22.1585 | -35.6544 | -35.7393
22.1646 | -35.655 | -35.7453
22.1707 | -35.6555 | -35.7521
22.1768 | -35.6561 | -35.7587
22.1829 | -35.6567 | -35.7651
22.189 | -35.6572 | -35.7716
22.1951 | -35.6578 | -35.7778
22.2012 | -35.6583 | -35.7843
22.2073 | -35.6589 | -35.7909
22.2134 | -35.6594 | -35.7973
22.2195 | -35.66 | -35.8036
22.2256 | -35.6606 | -35.8101
22.2317 | -35.6611 | -35.8164
22.2378 | -35.6617 | -35.8228
22.2439 | -35.6623 | -35.8293
22.25 | -35.6628 | -35.8358
22.2561 | -35.6634 | -35.8424
22.2622 | -35.664 | -35.8489
22.2683 | -35.6645 | -35.8549
22.2744 | -35.665 | -35.8613
22.2806 | -35.6656 | -35.8677
22.2867 | -35.6662 | -35.8741
22.2928 | -35.6667 | -35.8805
22.2989 | -35.6673 | -35.887
22.305 | -35.6678 | -35.8932
22.3111 | -35.6684 | -35.8998
22.3172 | -35.6689 | -35.9063
22.3233 | -35.6695 | -35.9127
22.3294 | -35.6701 | -35.9191
22.3355 | -35.6707 | -35.9253
22.3416 | -35.6712 | -35.9317
22.3477 | -35.6718 | -35.9382
22.3538 | -35.6724 | -35.9446
22.3599 | -35.6729 | -35.9508
22.366 | -35.6735 | -35.9573
22.3721 | -35.674 | -35.9635
22.3782 | -35.6746 | -35.97
22.3843 | -35.6751 | -35.9765
22.3904 | -35.6757 | -35.9832
22.3965 | -35.6763 | -35.9896
22.4026 | -35.6768 | -35.9959
22.4087 | -35.6773 | -36.0021
22.4148 | -35.6778 | -36.0085
22.4209 | -35.6784 | -36.015
22.4271 | -35.6789 | -36.0214
22.4332 | -35.6795 | -36.0275
22.4393 | -35.6801 | -36.0339
22.4454 | -35.6806 | -36.0398
22.4515 | -35.6811 | -36.0463
22.4576 | -35.6817 | -36.0527
22.4637 | -35.6822 | -36.0591
22.4698 | -35.6828 | -36.0655
22.4759 | -35.6834 | -36.0717
22.482 | -35.6839 | -36.0781
22.4881 | -35.6845 | -36.0845
22.4942 | -35.685 | -36.0909
22.5003 | -35.6856 | -36.0973
22.5064 | -35.6861 | -36.1036
22.5125 | -35.6866 | -36.1099
22.5186 | -35.6872 | -36.1162
22.5247 | -35.6877 | -36.1226
22.5308 | -35.6883 | -36.1293
22.5369 | -35.6889 | -36.1358
22.543 | -35.6894 | -36.1422
22.5491 | -35.69 | -36.1482
22.5552 | -35.6905 | -36.1547
22.5613 | -35.691 | -36.1612
22.5675 | -35.6916 | -36.1677
22.5736 | -35.6921 | -36.1742
22.5797 | -35.6927 | -36.1805
22.5858 | -35.6932 | -36.1866
22.5919 | -35.6937 | -36.1929
22.598 | -35.6943 | -36.1994
22.6041 | -35.6948 | -36.2057
22.6102 | -35.6954 | -36.2121
22.6163 | -35.6959 | -36.2183
22.6224 | -35.6965 | -36.2247
22.6285 | -35.697 | -36.2309
22.6346 | -35.6976 | -36.2371
22.6407 | -35.6981 | -36.2436
22.6468 | -35.6987 | -36.2499
22.6529 | -35.6992 | -36.256
22.659 | -35.6997 | -36.2623
22.6651 | -35.7003 | -36.2687
22.6712 | -35.7008 | -36.2749
22.6773 | -35.7014 | -36.2811
22.6834 | -35.7019 | -36.2877
22.6895 | -35.7025 | -36.2938
22.6956 | -35.703 | -36.3001
22.7017 | -35.7035 | -36.3065
22.7079 | -35.7041 | -36.313
22.714 | -35.7047 | -36.3194
22.7201 | -35.7052 | -36.3257
22.7262 | -35.7057 | -36.3319
22.7323 | -35.7063 | -36.3382
22.7384 | -35.7068 | -36.3446
22.7445 | -35.7074 | -36.3511
22.7506 | -35.7079 | -36.3574
22.7567 | -35.7085 | -36.3638
22.7628 | -35.709 | -36.3699
22.7689 | -35.7096 | -36.3761
22.775 | -35.7101 | -36.3826
22.7811 | -35.7107 | -36.3889
22.7872 | -35.7112 | -36.3952
22.7933 | -35.7117 | -36.4012
22.7994 | -35.7123 | -36.4075
22.8055 | -35.7128 | -36.4139
22.8116 | -35.7134 | -36.4201
22.8177 | -35.7139 | -36.4266
22.8238 | -35.7144 | -36.433
22.8299 | -35.715 | -36.439
22.836 | -35.7155 | -36.4455
22.8421 | -35.7161 | -36.4518
22.8482 | -35.7166 | -36.458
22.8544 | -35.7171 | -36.4645
22.8605 | -35.7177 | -36.4708
22.8666 | -35.7182 | -36.4771
22.8727 | -35.7188 | -36.4834
22.8788 | -35.7193 | -36.4897
22.8849 | -35.7199 | -36.4959
22.891 | -35.7204 | -36.5023
22.8971 | -35.721 | -36.5086
22.9032 | -35.7215 | -36.5146
22.9093 | -35.722 | -36.5209
22.9154 | -35.7226 | -36.5272
22.9215 | -35.7231 | -36.5335
22.9276 | -35.7236 | -36.5397
22.9337 | -35.7241 | -36.5457
22.9398 | -35.7247 | -36.552
22.9459 | -35.7252 | -36.5583
22.952 | -35.7258 | -36.5646
22.9581 | -35.7263 | -36.571
22.9642 | -35.7268 | -36.5774
22.9703 | -35.7274 | -36.5834
22.9764 | -35.7279 | -36.5897
22.9825 | -35.7285 | -36.5959
22.9886 | -35.729 | -36.6022
22.9948 | -35.7295 | -36.6085
23.0009 | -35.73 | -36.6148
23.007 | -35.7306 | -36.6211
23.0131 | -35.7311 | -36.6273
23.0192 | -35.7317 | -36.6333
23.0253 | -35.7322 | -36.6397
23.0314 | -35.7328 | -36.6459
23.0375 | -35.7333 | -36.6522
23.0436 | -35.7338 | -36.6584
23.0497 | -35.7344 | -36.6647
23.0558 | -35.735 | -36.6709
23.0619 | -35.7355 | -36.6773
23.068 | -35.7361 | -36.6836
23.0741 | -35.7366 | -36.6894
23.0802 | -35.7371 | -36.6957
23.0863 | -35.7376 | -36.702
23.0924 | -35.7381 | -36.7083
23.0985 | -35.7386 | -36.7145
23.1046 | -35.7392 | -36.7208
23.1107 | -35.7397 | -36.727
23.1168 | -35.7402 | -36.7333
23.1229 | -35.7407 | -36.7395
23.129 | -35.7413 | -36.746
23.1351 | -35.7418 | -36.7522
23.1413 | -35.7423 | -36.7585
23.1474 | -35.7429 | -36.7646
23.1535 | -35.7434 | -36.7708
23.1596 | -35.7439 | -36.7769
23.1657 | -35.7444 | -36.783
23.1718 | -35.745 | -36.7893
23.1779 | -35.7455 | -36.7957
23.184 | -35.746 | -36.8016
23.1901 | -35.7466 | -36.8079
23.1962 | -35.7471 | -36.8141
23.2023 | -35.7476 | -36.8202
23.2084 | -35.7482 | -36.8264
23.2145 | -35.7487 | -36.8325
23.2206 | -35.7493 | -36.8387
23.2267 | -35.7498 | -36.845
23.2328 | -35.7503 | -36.8512
23.2389 | -35.7508 | -36.8575
23.245 | -35.7514 | -36.8637
23.2511 | -35.7519 | -36.8698
23.2572 | -35.7525 | -36.8759
23.2633 | -35.753 | -36.8821
23.2694 | -35.7535 | -36.8884
23.2755 | -35.7541 | -36.8948
23.2817 | -35.7546 | -36.9009
23.2878 | -35.7551 | -36.907
23.2939 | -35.7556 | -36.9133
23.3 | -35.7562 | -36.9197
23.3061 | -35.7567 | -36.9259
23.3122 | -35.7572 | -36.932
23.3183 | -35.7578 | -36.9382
23.3244 | -35.7583 | -36.9444
23.3305 | -35.7588 | -36.9506
23.3366 | -35.7593 | -36.957
23.3427 | -35.7598 | -36.9631
23.3488 | -35.7604 | -36.9693
23.3549 | -35.7609 | -36.9756
23.361 | -35.7615 | -36.9816
23.3671 | -35.762 | -36.9879
23.3732 | -35.7626 | -36.9939
23.3793 | -35.7631 | -37
23.3854 | -35.7636 | -37.0063
23.3915 | -35.7641 | -37.0123
23.3976 | -35.7647 | -37.0186
23.4037 | -35.7652 | -37.0247
23.4098 | -35.7657 | -37.0309
23.4159 | -35.7663 | -37.037
23.422 | -35.7668 | -37.0431
23.4282 | -35.7673 | -37.049
23.4343 | -35.7679 | -37.0552
23.4404 | -35.7684 | -37.0616
23.4465 | -35.7689 | -37.0677
23.4526 | -35.7694 | -37.074
23.4587 | -35.7699 | -37.0802
23.4648 | -35.7704 | -37.086
23.4709 | -35.7709 | -37.0921
23.477 | -35.7714 | -37.0985
23.4831 | -35.7719 | -37.1046
23.4892 | -35.7724 | -37.1109
23.4953 | -35.773 | -37.117
23.5014 | -35.7735 | -37.123
23.5075 | -35.774 | -37.1292
23.5136 | -35.7745 | -37.1353
23.5197 | -35.7751 | -37.1415
23.5258 | -35.7756 | -37.1475
23.5319 | -35.7761 | -37.1535
23.538 | -35.7766 | -37.1597
23.5441 | -35.7771 | -37.1658
23.5502 | -35.7776 | -37.1721
23.5563 | -35.7782 | -37.1781
23.5624 | -35.7787 | -37.1844
23.5686 | -35.7792 | -37.1905
23.5747 | -35.7797 | -37.1966
23.5808 | -35.7803 | -37.2027
23.5869 | -35.7808 | -37.2088
23.593 | -35.7813 | -37.2149
23.5991 | -35.7818 | -37.2212
23.6052 | -35.7824 | -37.2271
23.6113 | -35.7829 | -37.2332
23.6174 | -35.7834 | -37.2393
23.6235 | -35.7839 | -37.2455
23.6296 | -35.7844 | -37.2519
23.6357 | -35.7849 | -37.2582
23.6418 | -35.7855 | -37.2641
23.6479 | -35.786 | -37.2704
23.654 | -35.7865 | -37.2765
23.6601 | -35.787 | -37.2826
23.6662 | -35.7875 | -37.2887
23.6723 | -35.788 | -37.2947
23.6784 | -35.7886 | -37.3008
23.6845 | -35.7891 | -37.307
23.6906 | -35.7896 | -37.3132
23.6967 | -35.7901 | -37.3193
23.7028 | -35.7906 | -37.3255
23.7089 | -35.7911 | -37.3317
23.7151 | -35.7916 | -37.3378
23.7212 | -35.7922 | -37.344
23.7273 | -35.7927 | -37.3502
23.7334 | -35.7932 | -37.3563
23.7395 | -35.7937 | -37.3624
23.7456 | -35.7942 | -37.3683
23.7517 | -35.7947 | -37.3745
23.7578 | -35.7952 | -37.3806
23.7639 | -35.7957 | -37.3867
23.77 | -35.7963 | -37.3928
23.7761 | -35.7968 | -37.3989
23.7822 | -35.7973 | -37.405
23.7883 | -35.7978 | -37.4111
23.7944 | -35.7983 | -37.4174
23.8005 | -35.7988 | -37.4235
23.8066 | -35.7994 | -37.4297
23.8127 | -35.7999 | -37.4358
23.8188 | -35.8004 | -37.442
23.8249 | -35.8009 | -37.4483
23.831 | -35.8014 | -37.4546
23.8371 | -35.8019 | -37.4605
23.8432 | -35.8025 | -37.4666
23.8493 | -35.803 | -37.4725
23.8555 | -35.8035 | -37.4787
23.8616 | -35.804 | -37.4849
23.8677 | -35.8045 | -37.491
23.8738 | -35.805 | -37.4971
23.8799 | -35.8056 | -37.5032
23.886 | -35.806 | -37.5091
23.8921 | -35.8065 | -37.5154
23.8982 | -35.8071 | -37.5215
23.9043 | -35.8076 | -37.5278
23.9104 | -35.8081 | -37.5339
23.9165 | -35.8086 | -37.5398
23.9226 | -35.8091 | -37.5456
23.9287 | -35.8096 | -37.5518
23.9348 | -35.8101 | -37.5577
23.9409 | -35.8106 | -37.5639
23.947 | -35.8111 | -37.57
23.9531 | -35.8116 | -37.5757
23.9592 | -35.8121 | -37.5819
23.9653 | -35.8126 | -37.5881
23.9714 | -35.8131 | -37.594
23.9775 | -35.8136 | -37.6002
23.9836 | -35.8141 | -37.6064
23.9897 | -35.8147 | -37.6124
23.9958 | -35.8152 | -37.6186
24.002 | -35.8157 | -37.6246
24.0081 | -35.8162 | -37.6305
24.0142 | -35.8167 | -37.6366
24.0203 | -35.8172 | -37.6425
24.0264 | -35.8177 | -37.6486
24.0325 | -35.8182 | -37.6547
24.0386 | -35.8187 | -37.6608
24.0447 | -35.8192 | -37.6668
24.0508 | -35.8197 | -37.6728
24.0569 | -35.8203 | -37.6791
24.063 | -35.8207 | -37.685
24.0691 | -35.8213 | -37.6909
24.0752 | -35.8218 | -37.6971
24.0813 | -35.8223 | -37.7031
24.0874 | -35.8228 | -37.7092
24.0935 | -35.8233 | -37.7153
24.0996 | -35.8238 | -37.7211
24.1057 | -35.8243 | -37.7271
24.1118 | -35.8248 | -37.7333
24.1179 | -35.8252 | -37.7396
24.124 | -35.8258 | -37.7455
24.1301 | -35.8263 | -37.7514
24.1362 | -35.8268 | -37.7574
24.1424 | -35.8273 | -37.7634
24.1485 | -35.8278 | -37.7693
24.1546 | -35.8283 | -37.7754
24.1607 | -35.8288 | -37.7815
24.1668 | -35.8293 | -37.7873
24.1729 | -35.8298 | -37.7933
24.179 | -35.8303 | -37.7993
24.1851 | -35.8308 | -37.8054
24.1912 | -35.8313 | -37.8116
24.1973 | -35.8318 | -37.8177
24.2034 | -35.8323 | -37.8235
24.2095 | -35.8328 | -37.8297
24.2156 | -35.8333 | -37.8359
24.2217 | -35.8338 | -37.842
24.2278 | -35.8344 | -37.848
24.2339 | -35.8349 | -37.854
24.24 | -35.8354 | -37.86
24.2461 | -35.8359 | -37.866
24.2522 | -35.8364 | -37.872
24.2583 | -35.8369 | -37.8783
24.2644 | -35.8374 | -37.8845
24.2705 | -35.8379 | -37.8905
24.2766 | -35.8384 | -37.8964
24.2827 | -35.8389 | -37.9022
24.2889 | -35.8394 | -37.9083
24.295 | -35.8399 | -37.9145
24.3011 | -35.8404 | -37.9205
24.3072 | -35.8409 | -37.9264
24.3133 | -35.8414 | -37.9323
24.3194 | -35.8419 | -37.9385
24.3255 | -35.8424 | -37.9447
24.3316 | -35.8429 | -37.9507
24.3377 | -35.8434 | -37.9567
24.3438 | -35.8438 | -37.9627
24.3499 | -35.8443 | -37.9687
24.356 | -35.8447 | -37.9747
24.3621 | -35.8452 | -37.9808
24.3682 | -35.8457 | -37.9867
24.3743 | -35.8462 | -37.993
24.3804 | -35.8467 | -37.9986
24.3865 | -35.8472 | -38.0047
24.3926 | -35.8476 | -38.0108
24.3987 | -35.8481 | -38.0169
24.4048 | -35.8486 | -38.0228
24.4109 | -35.8491 | -38.0286
24.417 | -35.8496 | -38.0345
24.4231 | -35.8501 | -38.0406
24.4293 | -35.8506 | -38.0465
24.4354 | -35.8511 | -38.0524
24.4415 | -35.8516 | -38.0582
24.4476 | -35.8521 | -38.0638
24.4537 | -35.8526 | -38.0697
24.4598 | -35.8532 | -38.0758
24.4659 | -35.8537 | -38.0819
24.472 | -35.8542 | -38.0879
24.4781 | -35.8547 | -38.094
24.4842 | -35.8551 | -38.0997
24.4903 | -35.8556 | -38.1056
24.4964 | -35.8561 | -38.1116
24.5025 | -35.8566 | -38.1177
24.5086 | -35.8571 | -38.1236
24.5147 | -35.8576 | -38.1295
24.5208 | -35.8581 | -38.1351
24.5269 | -35.8586 | -38.1409
24.533 | -35.8591 | -38.1467
24.5391 | -35.8596 | -38.1528
24.5452 | -35.8601 | -38.1588
24.5513 | -35.8606 | -38.1647
24.5574 | -35.8611 | -38.1708
24.5635 | -35.8616 | -38.1766
24.5696 | -35.8621 | -38.1825
24.5758 | -35.8626 | -38.1885
24.5819 | -35.8631 | -38.1944
24.588 | -35.8636 | -38.2002
24.5941 | -35.8641 | -38.2061
24.6002 | -35.8646 | -38.2121
24.6063 | -35.8651 | -38.2181
24.6124 | -35.8656 | -38.2243
24.6185 | -35.8661 | -38.2304
24.6246 | -35.8665 | -38.2363
24.6307 | -35.867 | -38.2424
24.6368 | -35.8675 | -38.2483
24.6429 | -35.868 | -38.2541
24.649 | -35.8685 | -38.26
24.6551 | -35.869 | -38.2659
24.6612 | -35.8695 | -38.2716
24.6673 | -35.87 | -38.2777
24.6734 | -35.8705 | -38.2837
24.6795 | -35.871 | -38.2897
24.6856 | -35.8715 | -38.2958
24.6917 | -35.872 | -38.3016
24.6978 | -35.8725 | -38.3074
24.7039 | -35.873 | -38.3133
24.71 | -35.8735 | -38.3194
24.7162 | -35.874 | -38.3255
24.7223 | -35.8744 | -38.3313
24.7284 | -35.8749 | -38.337
24.7345 | -35.8754 | -38.3429
24.7406 | -35.8759 | -38.3489
24.7467 | -35.8764 | -38.3549
24.7528 | -35.8769 | -38.3607
24.7589 | -35.8774 | -38.3666
24.765 | -35.8779 | -38.3722
24.7711 | -35.8784 | -38.3781
24.7772 | -35.8789 | -38.384
24.7833 | -35.8793 | -38.3899
24.7894 | -35.8798 | -38.3959
24.7955 | -35.8803 | -38.4018
24.8016 | -35.8808 | -38.4075
24.8077 | -35.8813 | -38.4136
24.8138 | -35.8818 | -38.4195
24.8199 | -35.8823 | -38.4255
24.826 | -35.8828 | -38.4314
24.8321 | -35.8833 | -38.4374
24.8382 | -35.8837 | -38.4431
24.8443 | -35.8842 | -38.449
24.8504 | -35.8847 | -38.4548
24.8565 | -35.8852 | -38.4607
24.8627 | -35.8857 | -38.4668
24.8688 | -35.8862 | -38.4725
24.8749 | -35.8867 | -38.4784
24.881 | -35.8871 | -38.4843
24.8871 | -35.8876 | -38.4902
24.8932 | -35.8882 | -38.496
24.8993 | -35.8886 | -38.5017
24.9054 | -35.8891 | -38.5075
24.9115 | -35.8896 | -38.5135
24.9176 | -35.8901 | -38.5195
24.9237 | -35.8906 | -38.5254
24.9298 | -35.8911 | -38.5313
24.9359 | -35.8916 | -38.5372
24.942 | -35.892 | -38.5429
24.9481 | -35.8925 | -38.5489
24.9542 | -35.893 | -38.5549
24.9603 | -35.8935 | -38.5608
24.9664 | -35.894 | -38.5667
24.9725 | -35.8945 | -38.5727
24.9786 | -35.8949 | -38.5784
24.9847 | -35.8954 | -38.5843
24.9908 | -35.8959 | -38.5902
24.9969 | -35.8963 | -38.5962
25.0031 | -35.8968 | -38.6021
25.0092 | -35.8973 | -38.6078
25.0153 | -35.8978 | -38.6135
25.0214 | -35.8982 | -38.6196
25.0275 | -35.8987 | -38.6253
25.0336 | -35.8992 | -38.6312
25.0397 | -35.8997 | -38.6373
25.0458 | -35.9002 | -38.643
25.0519 | -35.9007 | -38.6491
25.058 | -35.9012 | -38.6549
25.0641 | -35.9016 | -38.6607
25.0702 | -35.9021 | -38.6665
25.0763 | -35.9026 | -38.6723
25.0824 | -35.903 | -38.6781
25.0885 | -35.9035 | -38.6839
25.0946 | -35.904 | -38.6897
25.1007 | -35.9045 | -38.6956
25.1068 | -35.905 | -38.7014
25.1129 | -35.9054 | -38.7073
25.119 | -35.9059 | -38.7129
25.1251 | -35.9064 | -38.7188
25.1312 | -35.9069 | -38.7246
25.1373 | -35.9074 | -38.7306
25.1435 | -35.9079 | -38.7366
25.1496 | -35.9083 | -38.7423
25.1557 | -35.9088 | -38.7481
25.1618 | -35.9093 | -38.7541
25.1679 | -35.9098 | -38.7599
25.174 | -35.9102 | -38.7657
25.1801 | -35.9107 | -38.7717
25.1862 | -35.9112 | -38.7774
25.1923 | -35.9117 | -38.7832
25.1984 | -35.9122 | -38.789
25.2045 | -35.9127 | -38.7949
25.2106 | -35.9131 | -38.8008
25.2167 | -35.9136 | -38.8068
25.2228 | -35.914 | -38.8124
25.2289 | -35.9145 | -38.8182
25.235 | -35.915 | -38.824
25.2411 | -35.9155 | -38.8299
25.2472 | -35.9159 | -38.8358
25.2533 | -35.9164 | -38.8418
25.2594 | -35.9168 | -38.8473
25.2655 | -35.9173 | -38.8531
25.2716 | -35.9178 | -38.859
25.2777 | -35.9183 | -38.865
25.2838 | -35.9187 | -38.8709
25.29 | -35.9192 | -38.8766
25.2961 | -35.9197 | -38.8824
25.3022 | -35.9201 | -38.8881
25.3083 | -35.9206 | -38.8938
25.3144 | -35.9211 | -38.8997
25.3205 | -35.9215 | -38.9054
25.3266 | -35.922 | -38.9111
25.3327 | -35.9224 | -38.917
25.3388 | -35.9229 | -38.9228
25.3449 | -35.9234 | -38.9287
25.351 | -35.9238 | -38.9345
25.3571 | -35.9243 | -38.9404
25.3632 | -35.9248 | -38.9459
25.3693 | -35.9252 | -38.9518
25.3754 | -35.9257 | -38.9577
25.3815 | -35.9262 | -38.9635
25.3876 | -35.9267 | -38.9694
25.3937 | -35.9271 | -38.9751
25.3998 | -35.9276 | -38.9807
25.4059 | -35.9281 | -38.9867
25.412 | -35.9285 | -38.9925
25.4181 | -35.929 | -38.9984
25.4242 | -35.9295 | -39.0042
25.4304 | -35.93 | -39.01
25.4365 | -35.9304 | -39.0157
25.4426 | -35.9309 | -39.0215
25.4487 | -35.9313 | -39.0274
25.4548 | -35.9318 | -39.0332
25.4609 | -35.9323 | -39.039
25.467 | -35.9328 | -39.0446
25.4731 | -35.9332 | -39.0505
25.4792 | -35.9337 | -39.0562
25.4853 | -35.9342 | -39.0619
25.4914 | -35.9347 | -39.0677
25.4975 | -35.9351 | -39.0733
25.5036 | -35.9356 | -39.079
25.5097 | -35.9361 | -39.0847
25.5158 | -35.9365 | -39.0905
25.5219 | -35.937 | -39.0963
25.528 | -35.9375 | -39.1021
25.5341 | -35.938 | -39.1079
25.5402 | -35.9384 | -39.1136
25.5463 | -35.9389 | -39.1195
25.5524 | -35.9394 | -39.1254
25.5585 | -35.9398 | -39.131
25.5646 | -35.9403 | -39.1369
25.5707 | -35.9407 | -39.1427
25.5769 | -35.9412 | -39.1483
25.583 | -35.9417 | -39.1542
25.5891 | -35.9422 | -39.16
25.5952 | -35.9426 | -39.1658
25.6013 | -35.9431 | -39.1716
25.6074 | -35.9436 | -39.1771
25.6135 | -35.9441 | -39.1828
25.6196 | -35.9446 | -39.1888
25.6257 | -35.945 | -39.1946
25.6318 | -35.9455 | -39.2002
25.6379 | -35.9459 | -39.2059
25.644 | -35.9464 | -39.2114
25.6501 | -35.9469 | -39.217
25.6562 | -35.9473 | -39.2227
25.6623 | -35.9478 | -39.2285
25.6684 | -35.9483 | -39.2343
25.6745 | -35.9487 | -39.2401
25.6806 | -35.9492 | -39.2457
25.6867 | -35.9497 | -39.2515
25.6928 | -35.9501 | -39.2573
25.6989 | -35.9506 | -39.2631
25.705 | -35.951 | -39.269
25.7111 | -35.9515 | -39.2748
25.7173 | -35.9519 | -39.2804
25.7234 | -35.9524 | -39.2861
25.7295 | -35.9528 | -39.2919
25.7356 | -35.9533 | -39.2977
25.7417 | -35.9538 | -39.3034
25.7478 | -35.9542 | -39.3089
25.7539 | -35.9547 | -39.3145
25.76 | -35.9551 | -39.3202
25.7661 | -35.9556 | -39.3258
25.7722 | -35.9561 | -39.3315
25.7783 | -35.9565 | -39.3373
25.7844 | -35.957 | -39.3428
25.7905 | -35.9574 | -39.3486
25.7966 | -35.9579 | -39.3543
25.8027 | -35.9583 | -39.3601
25.8088 | -35.9588 | -39.366
25.8149 | -35.9592 | -39.3716
25.821 | -35.9596 | -39.3771
25.8271 | -35.9601 | -39.383
25.8332 | -35.9606 | -39.3888
25.8393 | -35.961 | -39.3945
25.8454 | -35.9615 | -39.4001
25.8515 | -35.962 | -39.4058
25.8576 | -35.9625 | -39.4113
25.8638 | -35.9629 | -39.4172
25.8699 | -35.9634 | -39.423
25.876 | -35.9638 | -39.4287
25.8821 | -35.9643 | -39.4345
25.8882 | -35.9647 | -39.4399
25.8943 | -35.9652 | -39.4457
25.9004 | -35.9656 | -39.4515
25.9065 | -35.9661 | -39.4572
25.9126 | -35.9666 | -39.4629
25.9187 | -35.9671 | -39.4686
25.9248 | -35.9675 | -39.4743
25.9309 | -35.968 | -39.48
25.937 | -35.9685 | -39.4859
25.9431 | -35.9689 | -39.4916
25.9492 | -35.9694 | -39.4972
25.9553 | -35.9698 | -39.5029
25.9614 | -35.9703 | -39.5083
25.9675 | -35.9707 | -39.5139
25.9736 | -35.9712 | -39.5197
25.9797 | -35.9717 | -39.5255
25.9858 | -35.9721 | -39.5311
25.9919 | -35.9726 | -39.537
25.998 | -35.973 | -39.5425
26.0042 | -35.9735 | -39.5482
26.0103 | -35.9739 | -39.5538
26.0164 | -35.9744 | -39.5594
26.0225 | -35.9748 | -39.565
26.0286 | -35.9753 | -39.5709
26.0347 | -35.9757 | -39.5764
26.0408 | -35.9762 | -39.5822
26.0469 | -35.9766 | -39.5877
26.053 | -35.9771 | -39.5934
26.0591 | -35.9776 | -39.5992
26.0652 | -35.978 | -39.6047
26.0713 | -35.9785 | -39.6105
26.0774 | -35.9789 | -39.6162
26.0835 | -35.9794 | -39.6219
26.0896 | -35.9798 | -39.6275
26.0957 | -35.9803 | -39.6331
26.1018 | -35.9807 | -39.6386
26.1079 | -35.9812 | -39.644
26.114 | -35.9816 | -39.6498
26.1201 | -35.982 | -39.6555
26.1262 | -35.9824 | -39.6612
26.1323 | -35.9829 | -39.6667
26.1384 | -35.9834 | -39.6723
26.1445 | -35.9838 | -39.678
26.1507 | -35.9843 | -39.6837
26.1568 | -35.9847 | -39.6893
26.1629 | -35.9852 | -39.695
26.169 | -35.9856 | -39.7006
26.1751 | -35.9861 | -39.7061
26.1812 | -35.9865 | -39.7117
26.1873 | -35.987 | -39.7173
26.1934 | -35.9875 | -39.7228
26.1995 | -35.9879 | -39.7284
26.2056 | -35.9883 | -39.7339
26.2117 | -35.9888 | -39.7396
26.2178 | -35.9892 | -39.7451
26.2239 | -35.9897 | -39.7506
26.23 | -35.9901 | -39.7564
26.2361 | -35.9906 | -39.762
26.2422 | -35.991 | -39.7675
26.2483 | -35.9915 | -39.773
26.2544 | -35.992 | -39.7784
26.2605 | -35.9924 | -39.7842
26.2666 | -35.9928 | -39.7898
26.2727 | -35.9933 | -39.7955
26.2788 | -35.9937 | -39.801
26.2849 | -35.9941 | -39.8066
26.2911 | -35.9946 | -39.8122
26.2972 | -35.995 | -39.8179
26.3033 | -35.9955 | -39.8235
26.3094 | -35.996 | -39.8292
26.3155 | -35.9964 | -39.8346
26.3216 | -35.9968 | -39.8403
26.3277 | -35.9973 | -39.8458
26.3338 | -35.9977 | -39.8512
26.3399 | -35.9982 | -39.8567
26.346 | -35.9986 | -39.8621
26.3521 | -35.999 | -39.8678
26.3582 | -35.9995 | -39.8734
26.3643 | -35.9999 | -39.879
26.3704 | -36.0004 | -39.8847
26.3765 | -36.0009 | -39.8903
26.3826 | -36.0013 | -39.8958
26.3887 | -36.0018 | -39.9014
26.3948 | -36.0022 | -39.907
26.4009 | -36.0026 | -39.9126
26.407 | -36.0031 | -39.9181
26.4131 | -36.0035 | -39.9237
26.4192 | -36.0039 | -39.9292
26.4253 | -36.0044 | -39.9347
26.4314 | -36.0049 | -39.9402
26.4376 | -36.0053 | -39.9459
26.4437 | -36.0057 | -39.9514
26.4498 | -36.0062 | -39.9571
26.4559 | -36.0066 | -39.9624
26.462 | -36.007 | -39.9681
26.4681 | -36.0075 | -39.9737
26.4742 | -36.0079 | -39.9794
26.4803 | -36.0084 | -39.985
26.4864 | -36.0088 | -39.9906
26.4925 | -36.0093 | -39.9962
26.4986 | -36.0097 | -40.0018
26.5047 | -36.0102 | -40.0075
26.5108 | -36.0107 | -40.013
26.5169 | -36.0111 | -40.0186
26.523 | -36.0115 | -40.0241
26.5291 | -36.0119 | -40.0295
26.5352 | -36.0123 | -40.0352
26.5413 | -36.0128 | -40.041
26.5474 | -36.0133 | -40.0465
26.5535 | -36.0137 | -40.0523
26.5596 | -36.0141 | -40.0578
26.5657 | -36.0145 | -40.0633
26.5718 | -36.015 | -40.069
26.578 | -36.0154 | -40.0746
26.5841 | -36.0159 | -40.0802
26.5902 | -36.0163 | -40.0857
26.5963 | -36.0167 | -40.0912
26.6024 | -36.0172 | -40.0968
26.6085 | -36.0176 | -40.1024
26.6146 | -36.018 | -40.1079
26.6207 | -36.0185 | -40.1136
26.6268 | -36.0189 | -40.1193
26.6329 | -36.0193 | -40.1247
26.639 | -36.0197 | -40.1304
26.6451 | -36.0202 | -40.136
26.6512 | -36.0206 | -40.1415
26.6573 | -36.021 | -40.1471
26.6634 | -36.0215 | -40.1524
26.6695 | -36.0219 | -40.1582
26.6756 | -36.0223 | -40.1638
26.6817 | -36.0227 | -40.1695
26.6878 | -36.0232 | -40.1752
26.6939 | -36.0236 | -40.1807
26.7 | -36.024 | -40.1861
26.7061 | -36.0245 | -40.1916
26.7122 | -36.0249 | -40.1973
26.7183 | -36.0253 | -40.2028
26.7245 | -36.0258 | -40.2084
26.7306 | -36.0262 | -40.214
26.7367 | -36.0267 | -40.2193
26.7428 | -36.0271 | -40.2248
26.7489 | -36.0275 | -40.2304
26.755 | -36.028 | -40.2358
26.7611 | -36.0284 | -40.2414
26.7672 | -36.0288 | -40.2469
26.7733 | -36.0293 | -40.2524
26.7794 | -36.0297 | -40.2579
26.7855 | -36.0302 | -40.2635
26.7916 | -36.0306 | -40.2689
26.7977 | -36.0311 | -40.2745
26.8038 | -36.0315 | -40.28
26.8099 | -36.032 | -40.2856
26.816 | -36.0324 | -40.2911
26.8221 | -36.0328 | -40.2965
26.8282 | -36.0332 | -40.3019
26.8343 | -36.0337 | -40.3075
26.8404 | -36.0341 | -40.3129
26.8465 | -36.0346 | -40.3184
26.8526 | -36.035 | -40.324
26.8587 | -36.0354 | -40.3295
26.8649 | -36.0358 | -40.335
26.871 | -36.0363 | -40.3406
26.8771 | -36.0367 | -40.346
26.8832 | -36.0371 | -40.3515
26.8893 | -36.0376 | -40.3571
26.8954 | -36.038 | -40.3627
26.9015 | -36.0384 | -40.3683
26.9076 | -36.0388 | -40.374
26.9137 | -36.0393 | -40.3795
26.9198 | -36.0397 | -40.3851
26.9259 | -36.0401 | -40.3905
26.932 | -36.0405 | -40.396
26.9381 | -36.041 | -40.4016
26.9442 | -36.0414 | -40.4068
26.9503 | -36.0418 | -40.4125
26.9564 | -36.0422 | -40.4181
26.9625 | -36.0427 | -40.4235
26.9686 | -36.0431 | -40.4291
26.9747 | -36.0436 | -40.4346
26.9808 | -36.044 | -40.4399
26.9869 | -36.0444 | -40.4454
26.993 | -36.0448 | -40.4509
26.9991 | -36.0453 | -40.4562
27.0052 | -36.0457 | -40.4619
27.0114 | -36.0461 | -40.4676
27.0175 | -36.0465 | -40.4729
27.0236 | -36.0469 | -40.4783
27.0297 | -36.0473 | -40.4839
27.0358 | -36.0478 | -40.4893
27.0419 | -36.0482 | -40.495
27.048 | -36.0486 | -40.5003
27.0541 | -36.049 | -40.5058
27.0602 | -36.0494 | -40.5115
27.0663 | -36.0499 | -40.517
27.0724 | -36.0503 | -40.5223
27.0785 | -36.0507 | -40.5279
27.0846 | -36.0511 | -40.5331
27.0907 | -36.0516 | -40.5387
27.0968 | -36.052 | -40.5442
27.1029 | -36.0524 | -40.5496
27.109 | -36.0529 | -40.555
27.1151 | -36.0533 | -40.5603
27.1212 | -36.0537 | -40.5653
27.1273 | -36.0541 | -40.5709
27.1334 | -36.0545 | -40.5766
27.1395 | -36.055 | -40.582
27.1456 | -36.0554 | -40.5876
27.1518 | -36.0558 | -40.5931
27.1579 | -36.0562 | -40.5984
27.164 | -36.0566 | -40.6038
27.1701 | -36.0571 | -40.6093
27.1762 | -36.0575 | -40.6148
27.1823 | -36.0579 | -40.6203
27.1884 | -36.0583 | -40.6258
27.1945 | -36.0587 | -40.6311
27.2006 | -36.0591 | -40.6365
27.2067 | -36.0595 | -40.642
27.2128 | -36.06 | -40.6475
27.2189 | -36.0604 | -40.6529
27.225 | -36.0608 | -40.6585
27.2311 | -36.0612 | -40.6638
27.2372 | -36.0616 | -40.6693
27.2433 | -36.0621 | -40.6748
27.2494 | -36.0625 | -40.6804
27.2555 | -36.0629 | -40.6859
27.2616 | -36.0634 | -40.6912
27.2677 | -36.0638 | -40.6966
27.2738 | -36.0642 | -40.702
27.2799 | -36.0646 | -40.7075
27.286 | -36.0651 | -40.713
27.2921 | -36.0655 | -40.7185
27.2983 | -36.0659 | -40.7238
27.3044 | -36.0663 | -40.7293
27.3105 | -36.0667 | -40.7348
27.3166 | -36.0671 | -40.7401
27.3227 | -36.0676 | -40.7455
27.3288 | -36.068 | -40.7508
27.3349 | -36.0684 | -40.7562
27.341 | -36.0688 | -40.7617
27.3471 | -36.0692 | -40.767
27.3532 | -36.0696 | -40.7723
27.3593 | -36.0701 | -40.7778
27.3654 | -36.0705 | -40.7833
27.3715 | -36.0709 | -40.7885
27.3776 | -36.0713 | -40.7941
27.3837 | -36.0718 | -40.7995
27.3898 | -36.0722 | -40.8048
27.3959 | -36.0726 | -40.8102
27.402 | -36.073 | -40.8155
27.4081 | -36.0734 | -40.8209
27.4142 | -36.0739 | -40.8264
27.4203 | -36.0743 | -40.8318
27.4264 | -36.0747 | -40.8372
27.4325 | -36.0751 | -40.8426
27.4387 | -36.0755 | -40.8479
27.4448 | -36.0759 | -40.8533
27.4509 | -36.0763 | -40.8588
27.457 | -36.0768 | -40.8643
27.4631 | -36.0772 | -40.8698
27.4692 | -36.0776 | -40.8751
27.4753 | -36.0779 | -40.8801
27.4814 | -36.0783 | -40.8855
27.4875 | -36.0787 | -40.8911
27.4936 | -36.0792 | -40.8965
27.4997 | -36.0796 | -40.9019
27.5058 | -36.08 | -40.9073
27.5119 | -36.0804 | -40.9125
27.518 | -36.0808 | -40.918
27.5241 | -36.0812 | -40.9234
27.5302 | -36.0816 | -40.9289
27.5363 | -36.082 | -40.9344
27.5424 | -36.0824 | -40.9396
27.5485 | -36.0828 | -40.945
27.5546 | -36.0833 | -40.9502
27.5607 | -36.0837 | -40.9556
27.5668 | -36.0841 | -40.9609
27.5729 | -36.0845 | -40.9662
27.5791 | -36.0849 | -40.9714
27.5852 | -36.0853 | -40.9767
27.5913 | -36.0857 | -40.9822
27.5974 | -36.0861 | -40.9877
27.6035 | -36.0865 | -40.9931
27.6096 | -36.0869 | -40.9983
27.6157 | -36.0873 | -41.0035
27.6218 | -36.0877 | -41.009
27.6279 | -36.0881 | -41.0145
27.634 | -36.0885 | -41.02
27.6401 | -36.0889 | -41.0255
27.6462 | -36.0894 | -41.0309
27.6523 | -36.0897 | -41.036
27.6584 | -36.0902 | -41.0414
27.6645 | -36.0906 | -41.0468
27.6706 | -36.0909 | -41.0522
27.6767 | -36.0914 | -41.0575
27.6828 | -36.0918 | -41.0627
27.6889 | -36.0922 | -41.068
27.695 | -36.0926 | -41.0734
27.7011 | -36.093 | -41.0788
27.7072 | -36.0934 | -41.0842
27.7133 | -36.0938 | -41.0896
27.7194 | -36.0942 | -41.0948
27.7256 | -36.0946 | -41.1001
27.7317 | -36.095 | -41.1054
27.7378 | -36.0954 | -41.1109
27.7439 | -36.0958 | -41.1163
27.75 | -36.0962 | -41.1217
27.7561 | -36.0966 | -41.1269
27.7622 | -36.097 | -41.1322
27.7683 | -36.0974 | -41.1377
27.7744 | -36.0978 | -41.1431
27.7805 | -36.0982 | -41.1485
27.7866 | -36.0987 | -41.1539
27.7927 | -36.099 | -41.1591
27.7988 | -36.0994 | -41.1644
27.8049 | -36.0998 | -41.1697
27.811 | -36.1002 | -41.175
27.8171 | -36.1007 | -41.1804
27.8232 | -36.1011 | -41.1855
27.8293 | -36.1015 | -41.1909
27.8354 | -36.1019 | -41.1963
27.8415 | -36.1023 | -41.2016
27.8476 | -36.1027 | -41.207
27.8537 | -36.1031 | -41.2123
27.8598 | -36.1035 | -41.2175
27.866 | -36.1039 | -41.2228
27.8721 | -36.1043 | -41.2281
27.8782 | -36.1047 | -41.2335
27.8843 | -36.1051 | -41.2389
27.8904 | -36.1055 | -41.2441
27.8965 | -36.1059 | -41.2494
27.9026 | -36.1062 | -41.2547
27.9087 | -36.1067 | -41.26
27.9148 | -36.1071 | -41.2652
27.9209 | -36.1074 | -41.2707
27.927 | -36.1078 | -41.2759
27.9331 | -36.1082 | -41.2813
27.9392 | -36.1087 | -41.2867
27.9453 | -36.1091 | -41.2919
27.9514 | -36.1094 | -41.2971
27.9575 | -36.1098 | -41.3025
27.9636 | -36.1102 | -41.3079
27.9697 | -36.1106 | -41.3131
27.9758 | -36.111 | -41.3185
27.9819 | -36.1114 | -41.3239
27.988 | -36.1118 | -41.3292
27.9941 | -36.1122 | -41.3346
28.0002 | -36.1125 | -41.3397
28.0063 | -36.1129 | -41.3451
28.0125 | -36.1133 | -41.3504
28.0186 | -36.1137 | -41.3559
28.0247 | -36.1141 | -41.361
28.0308 | -36.1145 | -41.3662
28.0369 | -36.1149 | -41.3712
28.043 | -36.1153 | -41.3768
28.0491 | -36.1157 | -41.3822
28.0552 | -36.116 | -41.3876
28.0613 | -36.1164 | -41.3931
28.0674 | -36.1168 | -41.3982
28.0735 | -36.1172 | -41.4033

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.8462 | -46.808
21.4908 | -36.0013 | -38.6045
16.4948 | -35.2516 | -30.888
12.6602 | -34.8056 | -24.1366
9.7171 | -34.5345 | -18.7458
7.4582 | -34.3457 | -14.326
5.7244 | -34.2152 | -10.7342
4.3936 | -34.1451 | -7.8182
3.3722 | -34.1125 | -5.5235
2.5883 | -34.1035 | -3.7826
1.9866 | -34.095 | -2.3901
1.5248 | -34.1033 | -1.2781
1.1703 | -34.1059 | -0.53577
1 | -34.1081 | -0.1712
0.89822 | -34.1011 | 0.12815
0.68941 | -34.087 | 0.76634
0.52919 | -34.0867 | 1.2579
0.40612 | -34.0918 | 1.8917
0.31171 | -34.0711 | 2.3704
0.23928 | -34.1037 | 3.0973
0.18358 | -34.1 | 3.5226
0.14091 | -34.1223 | 4.3884
0.10819 | -34.1415 | 5.527
0.083017 | -34.1686 | 6.9649
0.063705 | -34.2112 | 8.7854
0.048923 | -34.2549 | 11.2796
0.037479 | -34.3477 | 14.4781
0.028801 | -34.5051 | 18.7118
0.022078 | -34.8446 | 24.2685
0.017023 | -35.3442 | 30.7379
0.013018 | -36.1612 | 38.4629
0.010014 | -37.3122 | 46.4447

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.00382 | 0.94995
0.007639 | -28.702
0.011459 | -32.713
0.015279 | -37.2766
0.019099 | -39.9304
0.022918 | -41.1548
0.026738 | -43.4105
0.030558 | -44.5867
0.034377 | -46.3339
0.038197 | -48.5616
0.042017 | -50.0568
0.045837 | -51.2509
0.049656 | -51.9392
0.053476 | -52.8499
0.057296 | -54.2653
0.061115 | -55.8257
0.064935 | -55.9822
0.068755 | -56.5183
0.072574 | -57.5673
0.076394 | -58.2765
0.080214 | -58.4474
0.084034 | -58.5768
0.087853 | -59.9989
0.091673 | -61.6419
0.095493 | -61.4279
0.099312 | -61.2573
0.10313 | -62.524
0.10695 | -62.7332
0.11077 | -62.4829
0.11459 | -63.6327
0.11841 | -63.7991
0.12223 | -64.192
0.12605 | -63.5261
0.12987 | -64.1673
0.13369 | -64.8391
0.13751 | -65.7534
0.14133 | -65.8225
0.14515 | -66.1948
0.14897 | -66.4887
0.15279 | -67.0554
0.15661 | -66.7305
0.16043 | -67.4825
0.16425 | -67.6431
0.16807 | -67.4148
0.17189 | -68.0449
0.17571 | -68.4513
0.17953 | -68.7942
0.18335 | -69.2131
0.18717 | -68.8604
0.19098 | -69.2683
0.19481 | -69.6221
0.19862 | -70.5067
0.20245 | -70.3343
0.20626 | -70.8078
0.21008 | -70.9257
0.2139 | -71.1053
0.21772 | -72.2226
0.22154 | -71.1797
0.22536 | -71.9554
0.22918 | -72.1473
0.233 | -72.4083
0.23682 | -72.0976
0.24064 | -72.2199
0.24446 | -72.4589
0.24828 | -72.1131
0.2521 | -73.0531
0.25592 | -72.8923
0.25974 | -72.9856
0.26356 | -73.3863
0.26738 | -73.8183
0.2712 | -73.3782
0.27502 | -74.0145
0.27884 | -74.6307
0.28266 | -73.9715
0.28648 | -74.4175
0.2903 | -74.0979
0.29412 | -73.9071
0.29794 | -74.2889
0.30176 | -74.6844
0.30558 | -73.9912
0.3094 | -74.0903
0.31322 | -74.8287
0.31704 | -75.2048
0.32086 | -75.5889
0.32467 | -76.3758
0.32849 | -75.3454
0.33232 | -74.9254
0.33613 | -76.4071
0.33995 | -76.781
0.34377 | -76.4789
0.34759 | -76.884
0.35141 | -77.0519
0.35523 | -76.5125
0.35905 | -77.3858
0.36287 | -77.4612
0.36669 | -77.0946
0.37051 | -76.5589
0.37433 | -77.76
0.37815 | -77.5889
0.38197 | -76.5714
0.38579 | -76.7927
0.38961 | -77.4863
0.39343 | -77.5016
0.39725 | -77.6343
0.40107 | -77.9574
0.40489 | -77.6172
0.40871 | -77.3043
0.41253 | -77.4617
0.41635 | -78.7293
0.42017 | -78.2688
0.42399 | -78.7312
0.42781 | -78.9924
0.43163 | -79.1076
0.43545 | -78.2759
0.43927 | -78.5716
0.44309 | -78.5239
0.44691 | -78.2846
0.45073 | -79.4885
0.45454 | -78.7807
0.45837 | -78.1883
0.46219 | -79.7438
0.46601 | -80.4571
0.46982 | -79.336
0.47364 | -78.9546
0.47746 | -79.7667
0.48128 | -80.2471
0.4851 | -80.0487
0.48892 | -80.2208
0.49274 | -79.3443
0.49656 | -79.0734
0.50038 | -79.5558
0.5042 | -80.5856
0.50802 | -80.3602
0.51184 | -79.6396
0.51566 | -79.7787
0.51948 | -79.6883
0.5233 | -79.9314
0.52712 | -79.7815
0.53094 | -80.7242
0.53476 | -81.4304
0.53858 | -81.756
0.5424 | -80.8515
0.54622 | -80.0095
0.55004 | -80.0976
0.55386 | -80.3599
0.55768 | -80.7455
0.5615 | -80.9849
0.56532 | -80.9128
0.56914 | -81.081
0.57296 | -81.5768
0.57678 | -81.6557
0.5806 | -81.0895
0.58442 | -80.6406
0.58823 | -80.7096
0.59205 | -81.0961
0.59588 | -81.5729
0.59969 | -82.3527
0.60351 | -82.5242
0.60733 | -81.947
0.61115 | -81.6619
0.61497 | -81.6604
0.61879 | -81.5508
0.62261 | -81.556
0.62643 | -81.8553
0.63025 | -82.1481
0.63407 | -81.9952
0.63789 | -81.7532
0.64171 | -82.1925
0.64553 | -82.2191
0.64935 | -81.5306
0.65317 | -81.3108
0.65699 | -81.6218
0.66081 | -81.9254
0.66463 | -81.9393
0.66845 | -82.0608
0.67227 | -82.0969
0.67609 | -81.9923
0.67991 | -82.0252
0.68373 | -81.8284
0.68755 | -81.458
0.69137 | -81.5405
0.69519 | -82.0964
0.69901 | -82.517
0.70283 | -82.6032
0.70665 | -82.1575
0.71047 | -81.9534
0.71429 | -82.2411
0.7181 | -82.6142
0.72193 | -82.9796
0.72574 | -83.2475
0.72957 | -83.2621
0.73338 | -83.0855
0.7372 | -83.0455
0.74102 | -83.3402
0.74484 | -83.8349
0.74866 | -83.9972
0.75248 | -83.684
0.7563 | -83.3357
0.76012 | -83.2699
0.76394 | -83.5882
0.76776 | -83.919
0.77158 | -84.1407
0.7754 | -84.3002
0.77922 | -84.2172
0.78304 | -84.0394
0.78686 | -84
0.79068 | -84.3375
0.7945 | -84.6749
0.79832 | -84.5674
0.80214 | -84.2953
0.80596 | -84.3015
0.80978 | -84.1899
0.8136 | -83.7649
0.81742 | -83.5334
0.82124 | -83.7489
0.82506 | -84.1492
0.82888 | -84.2657
0.8327 | -84.4909
0.83652 | -85.069
0.84034 | -85.475
0.84416 | -85.4766
0.84798 | -85.4353
0.85179 | -85.1385
0.85562 | -84.7392
0.85943 | -84.5983
0.86325 | -84.5562
0.86707 | -84.3178
0.87089 | -84.1642
0.87471 | -84.1864
0.87853 | -84.1736
0.88235 | -84.6819
0.88617 | -85.3603
0.88999 | -85.6616
0.89381 | -85.9259
0.89763 | -85.8428
0.90145 | -85.345
0.90527 | -85.333
0.90909 | -85.8167
0.91291 | -85.8888
0.91673 | -85.6612
0.92055 | -85.6459
0.92437 | -85.579
0.92819 | -85.483
0.93201 | -85.5261
0.93583 | -85.7378
0.93965 | -85.6702
0.94347 | -85.3669
0.94729 | -85.7247
0.95111 | -86.3691
0.95493 | -86.3218
0.95875 | -86.1619
0.96257 | -86.3003
0.96639 | -86.4606
0.97021 | -86.413
0.97403 | -86.2012
0.97785 | -86.2165
0.98167 | -86.3508
0.98549 | -86.5296
0.9893 | -86.0357
0.99313 | -82.8554
0.99694 | -79.0734
1.0008 | -77.9074
1.0046 | -80.3326
1.0084 | -84.7463
1.0122 | -86.6247
1.016 | -86.4941
1.0199 | -86.7082
1.0237 | -86.8948
1.0275 | -86.5697
1.0313 | -86.2867
1.0351 | -86.3123
1.039 | -86.5088
1.0428 | -86.9938
1.0466 | -87.3904
1.0504 | -86.8881
1.0542 | -86.5677
1.0581 | -87.2354
1.0619 | -87.7179
1.0657 | -87.2862
1.0695 | -86.8319
1.0733 | -86.6879
1.0772 | -86.9214
1.081 | -87.2388
1.0848 | -87.2464
1.0886 | -87.3578
1.0924 | -87.6698
1.0963 | -87.9005
1.1001 | -87.7832
1.1039 | -87.7042
1.1077 | -87.5346
1.1115 | -87.339
1.1154 | -87.2285
1.1192 | -87.2126
1.123 | -87.3568
1.1268 | -87.6691
1.1306 | -87.8752
1.1345 | -87.7564
1.1383 | -87.5391
1.1421 | -87.5331
1.1459 | -87.8146
1.1497 | -88.1952
1.1536 | -88.4003
1.1574 | -88.4338
1.1612 | -88.4295
1.165 | -88.3452
1.1688 | -88.1637
1.1727 | -88.0832
1.1765 | -88.2832
1.1803 | -88.6778
1.1841 | -88.9087
1.1879 | -88.6764
1.1917 | -88.1708
1.1956 | -87.8046
1.1994 | -87.7143
1.2032 | -87.763
1.207 | -87.8593
1.2108 | -88.0419
1.2147 | -88.2238
1.2185 | -88.3405
1.2223 | -88.588
1.2261 | -88.9386
1.2299 | -88.9546
1.2338 | -88.564
1.2376 | -88.1466
1.2414 | -87.933
1.2452 | -88.0155
1.249 | -88.4319
1.2529 | -88.9962
1.2567 | -89.2627
1.2605 | -89.0038
1.2643 | -88.4749
1.2681 | -88.0933
1.272 | -88.055
1.2758 | -88.2075
1.2796 | -88.3807
1.2834 | -88.6215
1.2872 | -88.9002
1.2911 | -89.0682
1.2949 | -89.0884
1.2987 | -89.0903
1.3025 | -89.2172
1.3063 | -89.5203
1.3102 | -89.8067
1.314 | -89.7112
1.3178 | -89.3276
1.3216 | -89.0636
1.3254 | -89.0271
1.3293 | -89.0743
1.3331 | -88.9997
1.3369 | -88.7312
1.3407 | -88.414
1.3445 | -88.347
1.3484 | -88.7467
1.3522 | -89.393
1.356 | -89.6941
1.3598 | -89.4527
1.3636 | -89.0791
1.3675 | -89.0786
1.3713 | -89.6367
1.3751 | -90.3806
1.3789 | -90.6163
1.3827 | -90.3512
1.3866 | -90.0931
1.3904 | -89.8573
1.3942 | -89.4998
1.398 | -89.2663
1.4018 | -89.3706
1.4057 | -89.5833
1.4095 | -89.5052
1.4133 | -89.0922
1.4171 | -88.7029
1.4209 | -88.6947
1.4248 | -89.0385
1.4286 | -89.3259
1.4324 | -89.4101
1.4362 | -89.5504
1.44 | -89.824
1.4439 | -90.0857
1.4477 | -90.2233
1.4515 | -90.211
1.4553 | -90.2287
1.4591 | -90.4453
1.4629 | -90.6548
1.4668 | -90.5401
1.4706 | -90.2275
1.4744 | -90.117
1.4782 | -90.2891
1.482 | -90.3852
1.4859 | -90.1321
1.4897 | -89.7743
1.4935 | -89.6438
1.4973 | -89.8109
1.5011 | -90.1553
1.505 | -90.4813
1.5088 | -90.5619
1.5126 | -90.4065
1.5164 | -90.2675
1.5202 | -90.3145
1.5241 | -90.5695
1.5279 | -90.9367
1.5317 | -91.1904
1.5355 | -91.1444
1.5393 | -90.8342
1.5432 | -90.5586
1.547 | -90.4914
1.5508 | -90.5169
1.5546 | -90.5052
1.5584 | -90.5073
1.5623 | -90.5727
1.5661 | -90.621
1.5699 | -90.6114
1.5737 | -90.6622
1.5775 | -90.8539
1.5814 | -91.1657
1.5852 | -91.3967
1.589 | -91.3456
1.5928 | -91.0899
1.5966 | -90.7564
1.6005 | -90.5562
1.6043 | -90.7511
1.6081 | -91.2772
1.6119 | -91.6738
1.6157 | -91.5309
1.6196 | -91.0624
1.6234 | -90.7626
1.6272 | -90.8092
1.631 | -91.0418
1.6348 | -91.2089
1.6387 | -91.2839
1.6425 | -91.3221
1.6463 | -91.293
1.6501 | -91.131
1.6539 | -90.9506
1.6578 | -90.7689
1.6616 | -90.6645
1.6654 | -90.7113
1.6692 | -90.8806
1.673 | -91.0798
1.6769 | -91.218
1.6807 | -91.2532
1.6845 | -91.2683
1.6883 | -91.3551
1.6921 | -91.5051
1.696 | -91.6086
1.6998 | -91.5841
1.7036 | -91.4974
1.7074 | -91.416
1.7112 | -91.3109
1.715 | -91.1994
1.7189 | -91.1352
1.7227 | -91.1824
1.7265 | -91.3429
1.7303 | -91.5742
1.7341 | -91.8006
1.738 | -91.922
1.7418 | -91.8629
1.7456 | -91.6772
1.7494 | -91.4902
1.7532 | -91.4323
1.7571 | -91.5189
1.7609 | -91.6911
1.7647 | -91.87
1.7685 | -92.0243
1.7723 | -92.1263
1.7762 | -92.1208
1.78 | -91.9705
1.7838 | -91.7335
1.7876 | -91.5419
1.7914 | -91.465
1.7953 | -91.4532
1.7991 | -91.4154
1.8029 | -91.3295
1.8067 | -91.2133
1.8105 | -91.0866
1.8144 | -91.0011
1.8182 | -91.0281
1.822 | -91.205
1.8258 | -91.5162
1.8296 | -91.8966
1.8335 | -92.2321
1.8373 | -92.4153
1.8411 | -92.3949
1.8449 | -92.1919
1.8487 | -91.9681
1.8526 | -91.8835
1.8564 | -91.9093
1.8602 | -91.9433
1.864 | -91.9432
1.8678 | -91.9468
1.8717 | -91.9487
1.8755 | -91.9836
1.8793 | -92.093
1.8831 | -92.2238
1.8869 | -92.2597
1.8908 | -92.1704
1.8946 | -91.9982
1.8984 | -91.8166
1.9022 | -91.7015
1.906 | -91.7213
1.9099 | -91.8134
1.9137 | -91.8849
1.9175 | -91.9714
1.9213 | -92.177
1.9251 | -92.4378
1.929 | -92.5822
1.9328 | -92.5133
1.9366 | -92.313
1.9404 | -92.1394
1.9442 | -92.1387
1.9481 | -92.3482
1.9519 | -92.6401
1.9557 | -92.8188
1.9595 | -92.7954
1.9633 | -92.6297
1.9671 | -92.4312
1.971 | -92.2575
1.9748 | -92.1919
1.9786 | -92.2901
1.9824 | -91.8679
1.9862 | -90.1908
1.9901 | -87.8036
1.9939 | -85.6481
1.9977 | -84.3982
2.0015 | -84.2824
2.0053 | -85.3136
2.0092 | -87.3111
2.013 | -89.7311
2.0168 | -91.6916
2.0206 | -92.4779
2.0244 | -92.3549
2.0283 | -92.2274
2.0321 | -92.1893
2.0359 | -92.276
2.0397 | -92.4507
2.0435 | -92.6244
2.0474 | -92.7157
2.0512 | -92.7183
2.055 | -92.6656
2.0588 | -92.563
2.0626 | -92.465
2.0665 | -92.4254
2.0703 | -92.4254
2.0741 | -92.4321
2.0779 | -92.4281
2.0817 | -92.4503
2.0856 | -92.5624
2.0894 | -92.8195
2.0932 | -93.214
2.097 | -93.5834
2.1008 | -93.7173
2.1047 | -93.5968
2.1085 | -93.3736
2.1123 | -93.2041
2.1161 | -93.1023
2.1199 | -93.0209
2.1238 | -92.9294
2.1276 | -92.8304
2.1314 | -92.7659
2.1352 | -92.7962
2.139 | -92.9184
2.1429 | -93.065
2.1467 | -93.1786
2.1505 | -93.2636
2.1543 | -93.2698
2.1581 | -93.0968
2.162 | -92.7545
2.1658 | -92.4287
2.1696 | -92.3447
2.1734 | -92.603
2.1772 | -93.1143
2.1811 | -93.6222
2.1849 | -93.8843
2.1887 | -93.8933
2.1925 | -93.7862
2.1963 | -93.6514
2.2002 | -93.5227
2.204 | -93.4759
2.2078 | -93.453
2.2116 | -93.4886
2.2154 | -93.5899
2.2193 | -93.7553
2.2231 | -93.9719
2.2269 | -94.1909
2.2307 | -94.3366
2.2345 | -94.3491
2.2384 | -94.2041
2.2422 | -93.9508
2.246 | -93.6868
2.2498 | -93.4856
2.2536 | -93.3896
2.2574 | -93.3713
2.2613 | -93.3797
2.2651 | -93.3731
2.2689 | -93.3503
2.2727 | -93.3358
2.2765 | -93.3448
2.2804 | -93.3805
2.2842 | -93.4236
2.288 | -93.4597
2.2918 | -93.4758
2.2956 | -93.4879
2.2995 | -93.536
2.3033 | -93.6537
2.3071 | -93.852
2.3109 | -94.1216
2.3147 | -94.4394
2.3186 | -94.724
2.3224 | -94.8557
2.3262 | -94.7725
2.33 | -94.5084
2.3338 | -94.185
2.3377 | -93.9095
2.3415 | -93.7577
2.3453 | -93.7593
2.3491 | -93.8996
2.3529 | -94.1322
2.3568 | -94.3837
2.3606 | -94.5984
2.3644 | -94.7235
2.3682 | -94.731
2.372 | -94.6409
2.3759 | -94.4822
2.3797 | -94.2941
2.3835 | -94.097
2.3873 | -93.8845
2.3911 | -93.6596
2.395 | -93.4416
2.3988 | -93.2653
2.4026 | -93.1648
2.4064 | -93.1487
2.4102 | -93.2069
2.4141 | -93.3113
2.4179 | -93.4312
2.4217 | -93.5728
2.4255 | -93.7585
2.4293 | -94.0022
2.4332 | -94.3005
2.437 | -94.6025
2.4408 | -94.8529
2.4446 | -95.0211
2.4484 | -95.0923
2.4523 | -95.0908
2.4561 | -95.0533
2.4599 | -95.0181
2.4637 | -94.9538
2.4675 | -94.8363
2.4714 | -94.679
2.4752 | -94.5039
2.479 | -94.3698
2.4828 | -94.2545
2.4866 | -94.1253
2.4905 | -93.976
2.4943 | -93.8187
2.4981 | -93.7059
2.5019 | -93.6558
2.5057 | -93.7083
2.5095 | -93.8603
2.5134 | -94.065
2.5172 | -94.2643
2.521 | -94.3849
2.5248 | -94.4143
2.5286 | -94.3592
2.5325 | -94.2584
2.5363 | -94.1551
2.5401 | -94.0677
2.5439 | -94.0043
2.5477 | -93.9486
2.5516 | -93.9226
2.5554 | -93.9561
2.5592 | -94.0623
2.563 | -94.2298
2.5668 | -94.4183
2.5707 | -94.5761
2.5745 | -94.6528
2.5783 | -94.6291
2.5821 | -94.5329
2.5859 | -94.4104
2.5898 | -94.2873
2.5936 | -94.1676
2.5974 | -94.0765
2.6012 | -94.0279
2.605 | -94.0308
2.6089 | -94.0681
2.6127 | -94.1017
2.6165 | -94.0929
2.6203 | -94.0563
2.6241 | -94.047
2.628 | -94.1186
2.6318 | -94.2734
2.6356 | -94.4548
2.6394 | -94.5747
2.6432 | -94.6073
2.6471 | -94.5762
2.6509 | -94.5071
2.6547 | -94.4344
2.6585 | -94.403
2.6623 | -94.4368
2.6662 | -94.5518
2.67 | -94.7186
2.6738 | -94.8697
2.6776 | -94.9333
2.6814 | -94.8545
2.6853 | -94.6641
2.6891 | -94.4485
2.6929 | -94.2887
2.6967 | -94.2157
2.7005 | -94.2162
2.7044 | -94.2727
2.7082 | -94.3642
2.712 | -94.4931
2.7158 | -94.6565
2.7196 | -94.8428
2.7235 | -95.0422
2.7273 | -95.2121
2.7311 | -95.3194
2.7349 | -95.3579
2.7387 | -95.3378
2.7426 | -95.2806
2.7464 | -95.2394
2.7502 | -95.2712
2.754 | -95.4407
2.7578 | -95.5687
2.7616 | -95.6958
2.7655 | -95.7655
2.7693 | -95.7556
2.7731 | -95.6739
2.7769 | -95.5442
2.7807 | -95.3985
2.7846 | -95.2375
2.7884 | -95.0845
2.7922 | -94.9462
2.796 | -94.8366
2.7998 | -94.7563
2.8037 | -94.7283
2.8075 | -94.7328
2.8113 | -94.7606
2.8151 | -94.8074
2.8189 | -94.8365
2.8228 | -94.8434
2.8266 | -94.8225
2.8304 | -94.7745
2.8342 | -94.7239
2.838 | -94.7015
2.8419 | -94.7361
2.8457 | -94.8334
2.8495 | -94.9744
2.8533 | -95.1007
2.8571 | -95.1684
2.861 | -95.1568
2.8648 | -95.0791
2.8686 | -94.9831
2.8724 | -94.8537
2.8762 | -94.7132
2.8801 | -94.5624
2.8839 | -94.4305
2.8877 | -94.3411
2.8915 | -94.3011
2.8953 | -94.3041
2.8992 | -94.3374
2.903 | -94.3907
2.9068 | -94.4739
2.9106 | -94.582
2.9144 | -94.692
2.9183 | -94.7916
2.9221 | -94.8955
2.9259 | -94.9974
2.9297 | -95.1152
2.9335 | -95.242
2.9374 | -95.3725
2.9412 | -95.5008
2.945 | -95.6065
2.9488 | -95.6621
2.9526 | -95.6438
2.9565 | -95.5568
2.9603 | -95.4209
2.9641 | -95.2654
2.9679 | -95.1275
2.9717 | -95.0314
2.9756 | -94.9361
2.9794 | -94.8316
2.9832 | -94.6845
2.987 | -94.4225
2.9908 | -94.0734
2.9947 | -93.7272
2.9985 | -93.4907
3.0023 | -93.4395
3.0061 | -93.5583
3.0099 | -93.7813
3.0138 | -94.0209
3.0176 | -94.2111
3.0214 | -94.3448
3.0252 | -94.4623
3.029 | -94.5622
3.0328 | -94.6624
3.0367 | -94.8027
3.0405 | -94.961
3.0443 | -95.1038
3.0481 | -95.2096
3.0519 | -95.2581
3.0558 | -95.2705
3.0596 | -95.2555
3.0634 | -95.222
3.0672 | -95.19
3.071 | -95.1961
3.0749 | -95.2618
3.0787 | -95.3813
3.0825 | -95.5415
3.0863 | -95.6973
3.0901 | -95.807
3.094 | -95.8268
3.0978 | -95.6781
3.1016 | -95.2922
3.1054 | -94.7087
3.1092 | -94.0178
3.1131 | -93.3158
3.1169 | -92.7223
3.1207 | -92.3314
3.1245 | -92.2019
3.1283 | -92.3479
3.1322 | -92.7341
3.136 | -93.2904
3.1398 | -93.9045
3.1436 | -94.4614
3.1474 | -94.8939
3.1513 | -95.1856
3.1551 | -95.3397
3.1589 | -95.4157
3.1627 | -95.4681
3.1665 | -95.5128
3.1704 | -95.5487
3.1742 | -95.5874
3.178 | -95.6528
3.1818 | -95.7577
3.1856 | -95.8853
3.1895 | -96.0094
3.1933 | -96.105
3.1971 | -96.165
3.2009 | -96.2223
3.2047 | -96.2769
3.2086 | -96.3315
3.2124 | -96.3911
3.2162 | -96.4474
3.22 | -96.5292
3.2238 | -96.6305
3.2277 | -96.7254
3.2315 | -96.7898
3.2353 | -96.8135
3.2391 | -96.8054
3.2429 | -96.7772
3.2468 | -96.7146
3.2506 | -96.6213
3.2544 | -96.4974
3.2582 | -96.3507
3.262 | -96.2069
3.2659 | -96.1052
3.2697 | -96.0639
3.2735 | -96.1035
3.2773 | -96.2169
3.2811 | -96.3671
3.2849 | -96.5028
3.2888 | -96.5835
3.2926 | -96.599
3.2964 | -96.5641
3.3002 | -96.4884
3.304 | -96.4647
3.3079 | -96.4492
3.3117 | -96.4674
3.3155 | -96.5086
3.3193 | -96.5656
3.3231 | -96.6189
3.327 | -96.6541
3.3308 | -96.6767
3.3346 | -96.6959
3.3384 | -96.7333
3.3422 | -96.8017
3.3461 | -96.9076
3.3499 | -97.0367
3.3537 | -97.1665
3.3575 | -97.2568
3.3613 | -97.2856
3.3652 | -97.2593
3.369 | -97.1834
3.3728 | -97.0843
3.3766 | -96.9946
3.3804 | -96.9339
3.3843 | -96.8998
3.3881 | -96.8779
3.3919 | -96.872
3.3957 | -96.8675
3.3995 | -96.8524
3.4034 | -96.8364
3.4072 | -96.8301
3.411 | -96.8449
3.4148 | -96.8932
3.4186 | -96.9669
3.4225 | -97.0488
3.4263 | -97.1195
3.4301 | -97.1585
3.4339 | -97.1314
3.4377 | -97.049
3.4416 | -96.9376
3.4454 | -96.831
3.4492 | -96.743
3.453 | -96.703
3.4568 | -96.7147
3.4607 | -96.7653
3.4645 | -96.8465
3.4683 | -96.9442
3.4721 | -97.0372
3.4759 | -97.1243
3.4798 | -97.2008
3.4836 | -97.2741
3.4874 | -97.3385
3.4912 | -97.3993
3.495 | -97.451
3.4989 | -97.4672
3.5027 | -97.4404
3.5065 | -97.3701
3.5103 | -97.2656
3.5141 | -97.1613
3.518 | -97.0716
3.5218 | -97.0174
3.5256 | -97.0026
3.5294 | -97.0293
3.5332 | -97.0868
3.5371 | -97.1671
3.5409 | -97.263
3.5447 | -97.3592
3.5485 | -97.4291
3.5523 | -97.4685
3.5562 | -97.492
3.56 | -97.5014
3.5638 | -97.5016
3.5676 | -97.4909
3.5714 | -97.4771
3.5752 | -97.4725
3.5791 | -97.4565
3.5829 | -97.4322
3.5867 | -97.3924
3.5905 | -97.3257
3.5943 | -97.243
3.5982 | -97.1639
3.602 | -97.1027
3.6058 | -97.0676
3.6096 | -97.0541
3.6134 | -97.0629
3.6173 | -97.105
3.6211 | -97.1886
3.6249 | -97.3163
3.6287 | -97.4861
3.6325 | -97.6807
3.6364 | -97.8807
3.6402 | -98.0668
3.644 | -98.2207
3.6478 | -98.3233
3.6516 | -98.3589
3.6555 | -98.338
3.6593 | -98.2787
3.6631 | -98.191
3.6669 | -98.0961
3.6707 | -98.0248
3.6746 | -97.9656
3.6784 | -97.9199
3.6822 | -97.88
3.686 | -97.8307
3.6898 | -97.7743
3.6937 | -97.7098
3.6975 | -97.627
3.7013 | -97.5335
3.7051 | -97.4503
3.7089 | -97.3872
3.7128 | -97.3646
3.7166 | -97.4028
3.7204 | -97.4908
3.7242 | -97.6057
3.728 | -97.722
3.7319 | -97.8347
3.7357 | -97.9194
3.7395 | -97.9691
3.7433 | -97.9941
3.7471 | -98.0023
3.751 | -97.9954
3.7548 | -97.9678
3.7586 | -97.9388
3.7624 | -97.9056
3.7662 | -97.8785
3.7701 | -97.8529
3.7739 | -97.8281
3.7777 | -97.8225
3.7815 | -97.8205
3.7853 | -97.8146
3.7892 | -97.8066
3.793 | -97.8088
3.7968 | -97.8276
3.8006 | -97.8556
3.8044 | -97.8903
3.8083 | -97.9294
3.8121 | -97.9561
3.8159 | -97.9341
3.8197 | -97.8767
3.8235 | -97.8079
3.8273 | -97.7444
3.8312 | -97.7083
3.835 | -97.6967
3.8388 | -97.7149
3.8426 | -97.7588
3.8464 | -97.8072
3.8503 | -97.8312
3.8541 | -97.815
3.8579 | -97.8175
3.8617 | -97.8196
3.8655 | -97.8306
3.8694 | -97.852
3.8732 | -97.8889
3.877 | -97.9412
3.8808 | -97.982
3.8846 | -97.9958
3.8885 | -97.9768
3.8923 | -97.9207
3.8961 | -97.8349
3.8999 | -97.7292
3.9037 | -97.6169
3.9076 | -97.5285
3.9114 | -97.4858
3.9152 | -97.4855
3.919 | -97.5274
3.9228 | -97.6127
3.9267 | -97.7208
3.9305 | -97.8306
3.9343 | -97.9424
3.9381 | -98.0484
3.9419 | -98.1459
3.9458 | -98.2385
3.9496 | -98.3187
3.9534 | -98.3833
3.9572 | -98.4338
3.961 | -98.4797
3.9649 | -98.5084
3.9687 | -98.3899
3.9725 | -98.2073
3.9763 | -97.9396
3.9801 | -97.5542
3.984 | -97.0947
3.9878 | -96.6224
3.9916 | -96.2106
3.9954 | -95.9134
3.9992 | -95.7573
4.0031 | -95.7591
4.0069 | -95.9072
4.0107 | -96.172
4.0145 | -96.5066
4.0183 | -96.8578
4.0222 | -97.1799
4.026 | -97.4398
4.0298 | -97.6393
4.0336 | -97.8469
4.0374 | -97.9651
4.0413 | -98.0551
4.0451 | -98.1743
4.0489 | -98.3111
4.0527 | -98.4526
4.0565 | -98.5882
4.0604 | -98.7092
4.0642 | -98.8136
4.068 | -98.8984
4.0718 | -98.9544
4.0756 | -98.9859
4.0794 | -98.9981
4.0833 | -98.9957
4.0871 | -98.9772
4.0909 | -98.9618
4.0947 | -98.9465
4.0985 | -98.9174
4.1024 | -98.8835
4.1062 | -98.8308
4.11 | -98.7509
4.1138 | -98.6547
4.1176 | -98.5568
4.1215 | -98.4698
4.1253 | -98.3945
4.1291 | -98.3308
4.1329 | -98.2672
4.1367 | -98.2145
4.1406 | -98.1632
4.1444 | -98.1093
4.1482 | -98.0528
4.152 | -98.0051
4.1558 | -97.9676
4.1597 | -97.9555
4.1635 | -97.9722
4.1673 | -98.0083
4.1711 | -98.0496
4.1749 | -98.0879
4.1788 | -98.1137
4.1826 | -98.1185
4.1864 | -98.0997
4.1902 | -98.0671
4.194 | -98.0438
4.1979 | -98.0381
4.2017 | -98.0464
4.2055 | -98.0771
4.2093 | -98.1328
4.2131 | -98.206
4.217 | -98.275
4.2208 | -98.3331
4.2246 | -98.3837
4.2284 | -98.4166
4.2322 | -98.4482
4.2361 | -98.4663
4.2399 | -98.4557
4.2437 | -98.4232
4.2475 | -98.3761
4.2513 | -98.3295
4.2552 | -98.3041
4.259 | -98.2897
4.2628 | -98.3001
4.2666 | -98.3332
4.2704 | -98.37
4.2743 | -98.4066
4.2781 | -98.4419
4.2819 | -98.4774
4.2857 | -98.5105
4.2895 | -98.5351
4.2934 | -98.5671
4.2972 | -98.5999
4.301 | -98.6565
4.3048 | -98.7329
4.3086 | -98.809
4.3125 | -98.8795
4.3163 | -98.9275
4.3201 | -98.935
4.3239 | -98.9126
4.3277 | -98.8777
4.3316 | -98.8492
4.3354 | -98.8358
4.3392 | -98.8325
4.343 | -98.8405
4.3468 | -98.8622
4.3506 | -98.878
4.3545 | -98.8868
4.3583 | -98.8842
4.3621 | -98.859
4.3659 | -98.8127
4.3697 | -98.7519
4.3736 | -98.6942
4.3774 | -98.6368
4.3812 | -98.5868
4.385 | -98.5502
4.3888 | -98.535
4.3927 | -98.5429
4.3965 | -98.5829
4.4003 | -98.6615
4.4041 | -98.7787
4.4079 | -98.9071
4.4118 | -99.0164
4.4156 | -99.1085
4.4194 | -99.1863
4.4232 | -99.2286
4.427 | -99.2379
4.4309 | -99.2322
4.4347 | -99.2129
4.4385 | -99.1843
4.4423 | -99.145
4.4461 | -99.1088
4.45 | -99.0628
4.4538 | -99.0104
4.4576 | -98.9525
4.4614 | -98.8936
4.4652 | -98.8369
4.4691 | -98.7837
4.4729 | -98.7249
4.4767 | -98.6706
4.4805 | -98.6229
4.4843 | -98.5779
4.4882 | -98.5411
4.492 | -98.5161
4.4958 | -98.5137
4.4996 | -98.5313
4.5034 | -98.5509
4.5073 | -98.5723
4.5111 | -98.5983
4.5149 | -98.6294
4.5187 | -98.6682
4.5225 | -98.73
4.5264 | -98.8031
4.5302 | -98.8871
4.534 | -98.9771
4.5378 | -99.0659
4.5416 | -99.1468
4.5455 | -99.2123
4.5493 | -99.2564
4.5531 | -99.2911
4.5569 | -99.3273
4.5607 | -99.3636
4.5646 | -99.3902
4.5684 | -99.4141
4.5722 | -99.4304
4.576 | -99.4339
4.5798 | -99.419
4.5837 | -99.3849
4.5875 | -99.3206
4.5913 | -99.2331
4.5951 | -99.1364
4.5989 | -99.0499
4.6028 | -98.9672
4.6066 | -98.8957
4.6104 | -98.8517
4.6142 | -98.8394
4.618 | -98.8589
4.6218 | -98.9031
4.6257 | -98.961
4.6295 | -99.0287
4.6333 | -99.1133
4.6371 | -99.2073
4.6409 | -99.2924
4.6448 | -99.367
4.6486 | -99.4342
4.6524 | -99.4726
4.6562 | -99.4902
4.66 | -99.5024
4.6639 | -99.4974
4.6677 | -99.4707
4.6715 | -99.4336
4.6753 | -99.3981
4.6791 | -99.3591
4.683 | -99.3261
4.6868 | -99.2994
4.6906 | -99.2855
4.6944 | -99.2934
4.6982 | -99.3189
4.7021 | -99.3536
4.7059 | -99.3859
4.7097 | -99.4303
4.7135 | -99.4804
4.7173 | -99.5337
4.7212 | -99.5926
4.725 | -99.6512
4.7288 | -99.6981
4.7326 | -99.7298
4.7364 | -99.7509
4.7403 | -99.7688
4.7441 | -99.78
4.7479 | -99.7892
4.7517 | -99.8021
4.7555 | -99.8102
4.7594 | -99.8258
4.7632 | -99.8395
4.767 | -99.838
4.7708 | -99.8385
4.7746 | -99.8428
4.7785 | -99.8482
4.7823 | -99.8512
4.7861 | -99.8666
4.7899 | -99.8814
4.7937 | -99.8892
4.7976 | -99.8878
4.8014 | -99.8686
4.8052 | -99.8333
4.809 | -99.7774
4.8128 | -99.706
4.8167 | -99.6343
4.8205 | -99.5721
4.8243 | -99.5259
4.8281 | -99.4944
4.8319 | -99.483
4.8358 | -99.501
4.8396 | -99.5389
4.8434 | -99.5948
4.8472 | -99.6677
4.851 | -99.7462
4.8549 | -99.8125
4.8587 | -99.8606
4.8625 | -99.8932
4.8663 | -99.9052
4.8701 | -99.9133
4.8739 | -99.9142
4.8778 | -99.9208
4.8816 | -99.9417
4.8854 | -99.9751
4.8892 | -100.0188
4.893 | -100.0586
4.8969 | -100.0792
4.9007 | -100.0834
4.9045 | -100.0713
4.9083 | -100.0273
4.9121 | -99.9654
4.916 | -99.8817
4.9198 | -99.776
4.9236 | -99.6726
4.9274 | -99.5787
4.9312 | -99.5099
4.9351 | -99.476
4.9389 | -99.4787
4.9427 | -99.5087
4.9465 | -99.5527
4.9503 | -99.6021
4.9542 | -99.6436
4.958 | -99.5525
4.9618 | -99.1665
4.9656 | -98.7904
4.9694 | -98.3406
4.9733 | -97.7992
4.9771 | -97.1989
4.9809 | -96.5878
4.9847 | -96.0233
4.9885 | -95.5471
4.9924 | -95.1881
4.9962 | -94.9724
5 | -94.9107
5.0038 | -95.008
5.0076 | -95.2564
5.0115 | -95.6452
5.0153 | -96.1528
5.0191 | -96.7458
5.0229 | -97.3847
5.0267 | -98.0141
5.0306 | -98.593
5.0344 | -99.0841
5.0382 | -99.514
5.042 | -99.9773
5.0458 | -100.1061
5.0497 | -100.1282
5.0535 | -100.1351
5.0573 | -100.1319
5.0611 | -100.1174
5.0649 | -100.0944
5.0688 | -100.0653
5.0726 | -100.0444
5.0764 | -100.0244
5.0802 | -99.9965
5.084 | -99.9704
5.0879 | -99.9538
5.0917 | -99.9369
5.0955 | -99.9225
5.0993 | -99.9086
5.1031 | -99.8881
5.107 | -99.8677
5.1108 | -99.8435
5.1146 | -99.814
5.1184 | -99.7933
5.1222 | -99.786
5.1261 | -99.7911
5.1299 | -99.806
5.1337 | -99.8195
5.1375 | -99.8242
5.1413 | -99.8253
5.1451 | -99.8105
5.149 | -99.7828
5.1528 | -99.7429
5.1566 | -99.6996
5.1604 | -99.6721
5.1642 | -99.6462
5.1681 | -99.6272
5.1719 | -99.6187
5.1757 | -99.6103
5.1795 | -99.604
5.1833 | -99.6082
5.1872 | -99.6173
5.191 | -99.6441
5.1948 | -99.6921
5.1986 | -99.7528
5.2024 | -99.827
5.2063 | -99.9136
5.2101 | -100.0025
5.2139 | -100.0852
5.2177 | -100.1572
5.2215 | -100.2117
5.2254 | -100.2459
5.2292 | -100.2479
5.233 | -100.2284
5.2368 | -100.1952
5.2406 | -100.1404
5.2445 | -100.0919
5.2483 | -100.0535
5.2521 | -100.0274
5.2559 | -100.0105
5.2597 | -100.0015
5.2636 | -99.996
5.2674 | -100.001
5.2712 | -100.0114
5.275 | -100.0315
5.2788 | -100.0641
5.2827 | -100.1009
5.2865 | -100.1475
5.2903 | -100.2069
5.2941 | -100.2718
5.2979 | -100.3297
5.3018 | -100.3756
5.3056 | -100.4086
5.3094 | -100.4155
5.3132 | -100.3985
5.317 | -100.3601
5.3209 | -100.3063
5.3247 | -100.2365
5.3285 | -100.1494
5.3323 | -100.0578
5.3361 | -99.9684
5.34 | -99.876
5.3438 | -99.7763
5.3476 | -99.675
5.3514 | -99.5736
5.3552 | -99.4724
5.3591 | -99.3826
5.3629 | -99.3081
5.3667 | -99.2506
5.3705 | -99.2151
5.3743 | -99.2088
5.3782 | -99.236
5.382 | -99.2972
5.3858 | -99.3968
5.3896 | -99.5375
5.3934 | -99.703
5.3972 | -99.8797
5.4011 | -100.0643
5.4049 | -100.2408
5.4087 | -100.3947
5.4125 | -100.5053
5.4163 | -100.561
5.4202 | -100.5661
5.424 | -100.5274
5.4278 | -100.4497
5.4316 | -100.3615
5.4354 | -100.2727
5.4393 | -100.1865
5.4431 | -100.1155
5.4469 | -100.0614
5.4507 | -100.0238
5.4545 | -100.0013
5.4584 | -99.9909
5.4622 | -99.9847
5.466 | -99.9868
5.4698 | -99.9895
5.4736 | -99.9929
5.4775 | -100.0077
5.4813 | -100.0156
5.4851 | -100.013
5.4889 | -100.0004
5.4927 | -99.9798
5.4966 | -99.9495
5.5004 | -99.9086
5.5042 | -99.8605
5.508 | -99.807
5.5118 | -99.766
5.5157 | -99.727
5.5195 | -99.6985
5.5233 | -99.6764
5.5271 | -99.661
5.5309 | -99.6505
5.5348 | -99.6448
5.5386 | -99.6401
5.5424 | -99.6403
5.5462 | -99.6372
5.55 | -99.6356
5.5539 | -99.6352
5.5577 | -99.6358
5.5615 | -99.6298
5.5653 | -99.6259
5.5691 | -99.622
5.573 | -99.6198
5.5768 | -99.6172
5.5806 | -99.6118
5.5844 | -99.6071
5.5882 | -99.5962
5.5921 | -99.5769
5.5959 | -99.5477
5.5997 | -99.509
5.6035 | -99.4568
5.6073 | -99.3972
5.6112 | -99.336
5.615 | -99.2741
5.6188 | -99.2274
5.6226 | -99.1966
5.6264 | -99.1834
5.6303 | -99.1877
5.6341 | -99.2115
5.6379 | -99.2402
5.6417 | -99.2763
5.6455 | -99.3146
5.6494 | -99.348
5.6532 | -99.3731
5.657 | -99.3894
5.6608 | -99.395
5.6646 | -99.3903
5.6684 | -99.3868
5.6723 | -99.3825
5.6761 | -99.3774
5.6799 | -99.3649
5.6837 | -99.3421
5.6875 | -99.3115
5.6914 | -99.2836
5.6952 | -99.25
5.699 | -99.209
5.7028 | -99.1573
5.7066 | -99.0898
5.7105 | -99.0192
5.7143 | -98.9413
5.7181 | -98.8675
5.7219 | -98.7939
5.7257 | -98.7163
5.7296 | -98.6353
5.7334 | -98.5426
5.7372 | -98.4491
5.741 | -98.3552
5.7448 | -98.2609
5.7487 | -98.1721
5.7525 | -98.0936
5.7563 | -98.0268
5.7601 | -97.9762
5.7639 | -97.9462
5.7678 | -97.9349
5.7716 | -97.945
5.7754 | -97.9763
5.7792 | -98.0166
5.783 | -98.0752
5.7869 | -98.151
5.7907 | -98.2402
5.7945 | -98.3293
5.7983 | -98.4126
5.8021 | -98.4959
5.806 | -98.5649
5.8098 | -98.6266
5.8136 | -98.6835
5.8174 | -98.7362
5.8212 | -98.7862
5.8251 | -98.824
5.8289 | -98.852
5.8327 | -98.8834
5.8365 | -98.9135
5.8403 | -98.936
5.8442 | -98.9491
5.848 | -98.9484
5.8518 | -98.9359
5.8556 | -98.9105
5.8594 | -98.861
5.8633 | -98.7857
5.8671 | -98.6978
5.8709 | -98.595
5.8747 | -98.4767
5.8785 | -98.3531
5.8824 | -98.2374
5.8862 | -98.1352
5.89 | -98.0394
5.8938 | -97.9524
5.8976 | -97.8827
5.9015 | -97.829
5.9053 | -97.7888
5.9091 | -97.768
5.9129 | -97.7616
5.9167 | -97.7709
5.9206 | -97.7795
5.9244 | -97.7953
5.9282 | -97.8267
5.932 | -97.8655
5.9358 | -97.9172
5.9396 | -97.9766
5.9435 | -98.0412
5.9473 | -98.1066
5.9511 | -98.1726
5.9549 | -98.1809
5.9587 | -98.0918
5.9626 | -97.9968
5.9664 | -97.8579
5.9702 | -97.657
5.974 | -97.4125
5.9778 | -97.1434
5.9817 | -96.876
5.9855 | -96.6255
5.9893 | -96.4121
5.9931 | -96.2566
5.9969 | -96.1696
6.0008 | -96.161
6.0046 | -96.2344
6.0084 | -96.3877
6.0122 | -96.6112
6.016 | -96.8845
6.0199 | -97.1898
6.0237 | -97.5033
6.0275 | -97.8031
6.0313 | -98.0696
6.0351 | -98.2963
6.039 | -98.4785
6.0428 | -98.6511
6.0466 | -98.8168
6.0504 | -98.8864
6.0542 | -98.9402
6.0581 | -99.0245
6.0619 | -99.1005
6.0657 | -99.1811
6.0695 | -99.2707
6.0733 | -99.3653
6.0772 | -99.4583
6.081 | -99.5638
6.0848 | -99.6741
6.0886 | -99.7844
6.0924 | -99.8918
6.0963 | -99.9826
6.1001 | -100.0647
6.1039 | -100.1259
6.1077 | -100.1746
6.1115 | -100.2156
6.1154 | -100.2353
6.1192 | -100.24
6.123 | -100.2308
6.1268 | -100.2046
6.1306 | -100.1665
6.1345 | -100.1197
6.1383 | -100.0677
6.1421 | -100.0122
6.1459 | -99.9532
6.1497 | -99.8961
6.1536 | -99.8495
6.1574 | -99.8182
6.1612 | -99.7965
6.165 | -99.7826
6.1688 | -99.7717
6.1727 | -99.7696
6.1765 | -99.7571
6.1803 | -99.7389
6.1841 | -99.7087
6.1879 | -99.6544
6.1917 | -99.5771
6.1956 | -99.4497
6.1994 | -99.1008
6.2032 | -98.3757
6.207 | -97.6698
6.2108 | -96.9729
6.2147 | -96.239
6.2185 | -95.4837
6.2223 | -94.7351
6.2261 | -94.0248
6.2299 | -93.3819
6.2338 | -92.8255
6.2376 | -92.3726
6.2414 | -92.0347
6.2452 | -91.8197
6.249 | -91.7308
6.2529 | -91.7706
6.2567 | -91.9382
6.2605 | -92.2307
6.2643 | -92.6409
6.2681 | -93.1574
6.272 | -93.7669
6.2758 | -94.4495
6.2796 | -95.1819
6.2834 | -95.9346
6.2872 | -96.6735
6.2911 | -97.3754
6.2949 | -98.0859
6.2987 | -98.8665
6.3025 | -99.34
6.3063 | -99.5773
6.3102 | -99.7314
6.314 | -99.8511
6.3178 | -99.9527
6.3216 | -100.0372
6.3254 | -100.1093
6.3293 | -100.1756
6.3331 | -100.2298
6.3369 | -100.2773
6.3407 | -100.3186
6.3445 | -100.3542
6.3484 | -100.3791
6.3522 | -100.3986
6.356 | -100.4123
6.3598 | -100.4241
6.3636 | -100.4403
6.3675 | -100.4504
6.3713 | -100.454
6.3751 | -100.4614
6.3789 | -100.4701
6.3827 | -100.4787
6.3866 | -100.4891
6.3904 | -100.5002
6.3942 | -100.5038
6.398 | -100.5079
6.4018 | -100.5192
6.4057 | -100.5375
6.4095 | -100.5641
6.4133 | -100.6029
6.4171 | -100.6516
6.4209 | -100.7033
6.4248 | -100.765
6.4286 | -100.8407
6.4324 | -100.9238
6.4362 | -101.0096
6.44 | -101.0987
6.4439 | -101.1815
6.4477 | -101.2515
6.4515 | -101.3107
6.4553 | -101.3571
6.4591 | -101.392
6.4629 | -101.4216
6.4668 | -101.4396
6.4706 | -101.4495
6.4744 | -101.4563
6.4782 | -101.4556
6.482 | -101.4496
6.4859 | -101.4469
6.4897 | -101.4368
6.4935 | -101.4165
6.4973 | -101.3889
6.5011 | -101.3534
6.505 | -101.3157
6.5088 | -101.2685
6.5126 | -101.2168
6.5164 | -101.1699
6.5202 | -101.1244
6.5241 | -101.0747
6.5279 | -101.0359
6.5317 | -101.008
6.5355 | -100.9909
6.5393 | -100.9905
6.5432 | -101.003
6.547 | -101.0193
6.5508 | -101.0393
6.5546 | -101.0579
6.5584 | -101.0741
6.5623 | -101.0938
6.5661 | -101.1048
6.5699 | -101.1005
6.5737 | -101.0883
6.5775 | -101.0643
6.5814 | -101.0259
6.5852 | -100.9809
6.589 | -100.932
6.5928 | -100.8819
6.5966 | -100.8397
6.6005 | -100.8049
6.6043 | -100.7812
6.6081 | -100.7725
6.6119 | -100.8085
6.6157 | -100.822
6.6196 | -100.8514
6.6234 | -100.8987
6.6272 | -100.9513
6.631 | -101.0111
6.6348 | -101.0755
6.6387 | -101.1382
6.6425 | -101.1984
6.6463 | -101.2495
6.6501 | -101.2886
6.6539 | -101.3159
6.6578 | -101.3333
6.6616 | -101.3339
6.6654 | -101.3199
6.6692 | -101.3046
6.673 | -101.2836
6.6769 | -101.2602
6.6807 | -101.2437
6.6845 | -101.2297
6.6883 | -101.2234
6.6921 | -101.2259
6.696 | -101.237
6.6998 | -101.2586
6.7036 | -101.2957
6.7074 | -101.3415
6.7112 | -101.3969
6.715 | -101.4583
6.7189 | -101.5223
6.7227 | -101.5842
6.7265 | -101.6406
6.7303 | -101.6905
6.7341 | -101.7369
6.738 | -101.7763
6.7418 | -101.8047
6.7456 | -101.8253
6.7494 | -101.8417
6.7532 | -101.8557
6.7571 | -101.8611
6.7609 | -101.855
6.7647 | -101.8443
6.7685 | -101.831
6.7723 | -101.8112
6.7762 | -101.7795
6.78 | -101.743
6.7838 | -101.7008
6.7876 | -101.6671
6.7914 | -101.6319
6.7953 | -101.6011
6.7991 | -101.5769
6.8029 | -101.5534
6.8067 | -101.529
6.8105 | -101.5022
6.8144 | -101.4768
6.8182 | -101.4487
6.822 | -101.4159
6.8258 | -101.3782
6.8296 | -101.3432
6.8335 | -101.3076
6.8373 | -101.2697
6.8411 | -101.2394
6.8449 | -101.2179
6.8487 | -101.2019
6.8526 | -101.1896
6.8564 | -101.1813
6.8602 | -101.178
6.864 | -101.1802
6.8678 | -101.1835
6.8717 | -101.1907
6.8755 | -101.1928
6.8793 | -101.1979
6.8831 | -101.1985
6.8869 | -101.1957
6.8908 | -101.1881
6.8946 | -101.1782
6.8984 | -101.1677
6.9022 | -101.1638
6.906 | -101.1689
6.9099 | -101.1785
6.9137 | -101.1934
6.9175 | -101.2101
6.9213 | -101.2259
6.9251 | -101.2496
6.929 | -101.2767
6.9328 | -101.3035
6.9366 | -101.3378
6.9404 | -101.3792
6.9442 | -101.4222
6.9481 | -101.456
6.9519 | -101.4846
6.9557 | -101.519
6.9595 | -101.5406
6.9633 | -101.5517
6.9672 | -101.5491
6.971 | -101.5323
6.9748 | -101.5037
6.9786 | -101.4649
6.9824 | -101.4112
6.9862 | -101.3564
6.9901 | -101.3016
6.9939 | -101.2467
6.9977 | -101.199
7.0015 | -101.161
7.0053 | -101.123
7.0092 | -101.0932
7.013 | -101.0728
7.0168 | -101.0637
7.0206 | -101.0575
7.0244 | -101.0613
7.0283 | -101.0729
7.0321 | -101.0921
7.0359 | -101.1187
7.0397 | -101.1537
7.0435 | -101.1966
7.0474 | -101.2476
7.0512 | -101.3126
7.055 | -101.3859
7.0588 | -101.4542
7.0626 | -101.5198
7.0665 | -101.5767
7.0703 | -101.6287
7.0741 | -101.6651
7.0779 | -101.6833
7.0817 | -101.6978
7.0856 | -101.6984
7.0894 | -101.6823
7.0932 | -101.6576
7.097 | -101.6258
7.1008 | -101.5897
7.1047 | -101.552
7.1085 | -101.5157
7.1123 | -101.4831
7.1161 | -101.4558
7.1199 | -101.4261
7.1238 | -101.4029
7.1276 | -101.385
7.1314 | -101.3659
7.1352 | -101.3495
7.139 | -101.3319
7.1429 | -101.3129
7.1467 | -101.293
7.1505 | -101.2771
7.1543 | -101.2612
7.1581 | -101.2394
7.162 | -101.2339
7.1658 | -101.222
7.1696 | -101.2207
7.1734 | -101.2245
7.1772 | -101.2318
7.1811 | -101.2492
7.1849 | -101.2791
7.1887 | -101.3119
7.1925 | -101.3538
7.1963 | -101.4024
7.2002 | -101.4574
7.204 | -101.5196
7.2078 | -101.5817
7.2116 | -101.6386
7.2154 | -101.7032
7.2193 | -101.7612
7.2231 | -101.8101
7.2269 | -101.8576
7.2307 | -101.8957
7.2345 | -101.9208
7.2383 | -101.9368
7.2422 | -101.9337
7.246 | -101.9263
7.2498 | -101.9052
7.2536 | -101.8773
7.2574 | -101.8415
7.2613 | -101.8012
7.2651 | -101.7611
7.2689 | -101.7184
7.2727 | -101.683
7.2765 | -101.6576
7.2804 | -101.6363
7.2842 | -101.626
7.288 | -101.628
7.2918 | -101.6385
7.2956 | -101.6466
7.2995 | -101.6611
7.3033 | -101.6779
7.3071 | -101.6902
7.3109 | -101.7017
7.3147 | -101.7099
7.3186 | -101.7126
7.3224 | -101.709
7.3262 | -101.6999
7.33 | -101.688
7.3338 | -101.6736
7.3377 | -101.6588
7.3415 | -101.6437
7.3453 | -101.6281
7.3491 | -101.6173
7.3529 | -101.6174
7.3568 | -101.6157
7.3606 | -101.6178
7.3644 | -101.6233
7.3682 | -101.6338
7.372 | -101.6504
7.3759 | -101.6708
7.3797 | -101.6917
7.3835 | -101.7115
7.3873 | -101.7277
7.3911 | -101.7429
7.395 | -101.7482
7.3988 | -101.7485
7.4026 | -101.7543
7.4064 | -101.764
7.4102 | -101.7674
7.4141 | -101.769
7.4179 | -101.7737
7.4217 | -101.7755
7.4255 | -101.7816
7.4293 | -101.7914
7.4332 | -101.7997
7.437 | -101.7956
7.4408 | -101.7814
7.4446 | -101.7616
7.4484 | -101.7386
7.4523 | -101.7077
7.4561 | -101.6653
7.4599 | -101.6125
7.4637 | -101.5622
7.4675 | -101.5141
7.4714 | -101.4751
7.4752 | -101.4397
7.479 | -101.4117
7.4828 | -101.3938
7.4866 | -101.3838
7.4905 | -101.3841
7.4943 | -101.3927
7.4981 | -101.4178
7.5019 | -101.4528
7.5057 | -101.503
7.5095 | -101.5576
7.5134 | -101.6166
7.5172 | -101.6827
7.521 | -101.7583
7.5248 | -101.8366
7.5286 | -101.9101
7.5325 | -101.9804
7.5363 | -102.0428
7.5401 | -102.0975
7.5439 | -102.1452
7.5477 | -102.1829
7.5516 | -102.2102
7.5554 | -102.2307
7.5592 | -102.2401
7.563 | -102.2497
7.5668 | -102.2635
7.5707 | -102.2785
7.5745 | -102.2844
7.5783 | -102.2911
7.5821 | -102.2919
7.5859 | -102.2941
7.5898 | -102.2955
7.5936 | -102.2973
7.5974 | -102.2967
7.6012 | -102.2893
7.605 | -102.2726
7.6089 | -102.2419
7.6127 | -102.2116
7.6165 | -102.1709
7.6203 | -102.1306
7.6241 | -102.0865
7.628 | -102.035
7.6318 | -101.9773
7.6356 | -101.9119
7.6394 | -101.8413
7.6432 | -101.7607
7.6471 | -101.6758
7.6509 | -101.5833
7.6547 | -101.4852
7.6585 | -101.3898
7.6623 | -101.2926
7.6662 | -101.1899
7.67 | -101.0935
7.6738 | -101.0092
7.6776 | -100.9424
7.6814 | -100.8919
7.6853 | -100.8614
7.6891 | -100.8491
7.6929 | -100.8631
7.6967 | -100.897
7.7005 | -100.9476
7.7044 | -101.0126
7.7082 | -101.095
7.712 | -101.1957
7.7158 | -101.2779
7.7196 | -101.3639
7.7235 | -101.4545
7.7273 | -101.5404
7.7311 | -101.6219
7.7349 | -101.6952
7.7387 | -101.7617
7.7426 | -101.8257
7.7464 | -101.8843
7.7502 | -101.9396
7.754 | -101.9905
7.7578 | -102.0401
7.7617 | -102.0841
7.7655 | -102.1279
7.7693 | -102.1682
7.7731 | -102.2058
7.7769 | -102.2502
7.7807 | -102.2984
7.7846 | -102.3433
7.7884 | -102.3834
7.7922 | -102.4225
7.796 | -102.4649
7.7998 | -102.5066
7.8037 | -102.5485
7.8075 | -102.5821
7.8113 | -102.6109
7.8151 | -102.6344
7.8189 | -102.6494
7.8228 | -102.6534
7.8266 | -102.6542
7.8304 | -102.653
7.8342 | -102.6527
7.838 | -102.6518
7.8419 | -102.6471
7.8457 | -102.6415
7.8495 | -102.6408
7.8533 | -102.6495
7.8571 | -102.6624
7.861 | -102.6784
7.8648 | -102.6954
7.8686 | -102.7131
7.8724 | -102.7282
7.8762 | -102.7403
7.8801 | -102.7523
7.8839 | -102.7573
7.8877 | -102.7534
7.8915 | -102.7405
7.8953 | -102.721
7.8992 | -102.6983
7.903 | -102.6772
7.9068 | -102.6561
7.9106 | -102.632
7.9144 | -102.6029
7.9183 | -102.5767
7.9221 | -102.5569
7.9259 | -102.5365
7.9297 | -102.5182
7.9335 | -102.5017
7.9374 | -102.4895
7.9412 | -102.4631
7.945 | -102.4311
7.9488 | -102.4059
7.9526 | -102.3773
7.9565 | -102.3425
7.9603 | -102.2998
7.9641 | -102.252
7.9679 | -102.1952
7.9717 | -102.1264
7.9756 | -102.0569
7.9794 | -101.9854
7.9832 | -101.9094
7.987 | -101.8384
7.9908 | -101.7839
7.9947 | -101.7432
7.9985 | -101.7177
8.0023 | -101.7129
8.0061 | -101.7297
8.0099 | -101.7645
8.0138 | -101.8135
8.0176 | -101.8733
8.0214 | -101.9439
8.0252 | -102.0218
8.029 | -102.1005
8.0328 | -102.1826
8.0367 | -102.2615
8.0405 | -102.3326
8.0443 | -102.4027
8.0481 | -102.4674
8.0519 | -102.5218
8.0558 | -102.5687
8.0596 | -102.6173
8.0634 | -102.6557
8.0672 | -102.6732
8.071 | -102.6829
8.0749 | -102.6846
8.0787 | -102.6851
8.0825 | -102.6819
8.0863 | -102.6751
8.0901 | -102.6588
8.094 | -102.6387
8.0978 | -102.62
8.1016 | -102.611
8.1054 | -102.6011
8.1092 | -102.5801
8.1131 | -102.5645
8.1169 | -102.5552
8.1207 | -102.555
8.1245 | -102.5618
8.1283 | -102.5768
8.1322 | -102.594
8.136 | -102.6134
8.1398 | -102.6365
8.1436 | -102.658
8.1474 | -102.6827
8.1513 | -102.7038
8.1551 | -102.7167
8.1589 | -102.7268
8.1627 | -102.7351
8.1665 | -102.7395
8.1704 | -102.742
8.1742 | -102.7384
8.178 | -102.7317
8.1818 | -102.7189
8.1856 | -102.7005
8.1895 | -102.6771
8.1933 | -102.6553
8.1971 | -102.6315
8.2009 | -102.6008
8.2047 | -102.564
8.2086 | -102.535
8.2124 | -102.5005
8.2162 | -102.4685
8.22 | -102.4367
8.2238 | -102.4024
8.2277 | -102.379
8.2315 | -102.3583
8.2353 | -102.338
8.2391 | -102.3111
8.2429 | -102.2831
8.2468 | -102.2653
8.2506 | -102.2505
8.2544 | -102.2346
8.2582 | -102.2222
8.262 | -102.21
8.2659 | -102.1186
8.2697 | -102.0027
8.2735 | -101.8232
8.2773 | -101.6259
8.2811 | -101.4196
8.285 | -101.1986
8.2888 | -100.9524
8.2926 | -100.6829
8.2964 | -100.3969
8.3002 | -100.1038
8.304 | -99.8141
8.3079 | -99.5332
8.3117 | -99.275
8.3155 | -99.046
8.3193 | -98.8526
8.3231 | -98.7023
8.327 | -98.5993
8.3308 | -98.5435
8.3346 | -98.5412
8.3384 | -98.5885
8.3422 | -98.6875
8.3461 | -98.8374
8.3499 | -99.0301
8.3537 | -99.2618
8.3575 | -99.5297
8.3613 | -99.828
8.3652 | -100.1445
8.369 | -100.4631
8.3728 | -100.7816
8.3766 | -101.0938
8.3804 | -101.3846
8.3843 | -101.6497
8.3881 | -101.8978
8.3919 | -102.1311
8.3957 | -102.3436
8.3995 | -102.5481
8.4034 | -102.6897
8.4072 | -102.7655
8.411 | -102.804
8.4148 | -102.836
8.4186 | -102.8653
8.4225 | -102.8841
8.4263 | -102.8994
8.4301 | -102.9063
8.4339 | -102.9087
8.4377 | -102.9004
8.4416 | -102.8841
8.4454 | -102.8656
8.4492 | -102.8477
8.453 | -102.8265
8.4568 | -102.8056
8.4607 | -102.7827
8.4645 | -102.7648
8.4683 | -102.7487
8.4721 | -102.7364
8.4759 | -102.7265
8.4798 | -102.7145
8.4836 | -102.7041
8.4874 | -102.6957
8.4912 | -102.6881
8.495 | -102.6813
8.4989 | -102.6727
8.5027 | -102.6589
8.5065 | -102.6407
8.5103 | -102.6273
8.5141 | -102.6169
8.518 | -102.6037
8.5218 | -102.5927
8.5256 | -102.5825
8.5294 | -102.5725
8.5332 | -102.5705
8.5371 | -102.5749
8.5409 | -102.5808
8.5447 | -102.5861
8.5485 | -102.5974
8.5523 | -102.6064
8.5562 | -102.6168
8.56 | -102.6298
8.5638 | -102.6358
8.5676 | -102.6391
8.5714 | -102.6426
8.5752 | -102.642
8.5791 | -102.637
8.5829 | -102.6302
8.5867 | -102.6246
8.5905 | -102.6186
8.5943 | -102.6121
8.5982 | -102.6053
8.602 | -102.6029
8.6058 | -102.6052
8.6096 | -102.608
8.6134 | -102.614
8.6173 | -102.6243
8.6211 | -102.6407
8.6249 | -102.6581
8.6287 | -102.6722
8.6325 | -102.6875
8.6364 | -102.7007
8.6402 | -102.7144
8.644 | -102.73
8.6478 | -102.7399
8.6516 | -102.7495
8.6555 | -102.7606
8.6593 | -102.7654
8.6631 | -102.7636
8.6669 | -102.7618
8.6707 | -102.7643
8.6746 | -102.763
8.6784 | -102.7578
8.6822 | -102.7502
8.686 | -102.7405
8.6898 | -102.7333
8.6937 | -102.7186
8.6975 | -102.6954
8.7013 | -102.668
8.7051 | -102.6387
8.7089 | -102.6045
8.7128 | -102.5684
8.7166 | -102.527
8.7204 | -102.4858
8.7242 | -102.4423
8.728 | -102.3954
8.7319 | -102.3502
8.7357 | -102.3067
8.7395 | -102.2629
8.7433 | -102.2199
8.7471 | -102.1813
8.751 | -102.1489
8.7548 | -102.1226
8.7586 | -102.101
8.7624 | -102.0814
8.7662 | -102.0623
8.7701 | -102.0476
8.7739 | -102.0392
8.7777 | -102.0302
8.7815 | -102.0254
8.7853 | -102.0261
8.7892 | -102.0285
8.793 | -102.0326
8.7968 | -102.0346
8.8006 | -102.0381
8.8044 | -102.0439
8.8083 | -102.0518
8.8121 | -102.063
8.8159 | -102.0851
8.8197 | -102.099
8.8235 | -102.1087
8.8273 | -102.1186
8.8312 | -102.1273
8.835 | -102.1388
8.8388 | -102.1478
8.8426 | -102.1496
8.8464 | -102.1511
8.8503 | -102.1496
8.8541 | -102.1473
8.8579 | -102.1472
8.8617 | -102.144
8.8655 | -102.1353
8.8694 | -102.131
8.8732 | -102.1278
8.877 | -102.1231
8.8808 | -102.1225
8.8846 | -102.1256
8.8885 | -102.1327
8.8923 | -102.1466
8.8961 | -102.1641
8.8999 | -102.1832
8.9037 | -102.2062
8.9076 | -102.2288
8.9114 | -102.2501
8.9152 | -102.27
8.919 | -102.2901
8.9228 | -102.3069
8.9267 | -102.3196
8.9305 | -102.3199
8.9343 | -102.2774
8.9381 | -102.2381
8.9419 | -102.2038
8.9458 | -102.1614
8.9496 | -102.1156
8.9534 | -102.0648
8.9572 | -102.012
8.961 | -101.9572
8.9649 | -101.9024
8.9687 | -101.8437
8.9725 | -101.7889
8.9763 | -101.7438
8.9801 | -101.7069
8.984 | -101.6761
8.9878 | -101.6538
8.9916 | -101.6404
8.9954 | -101.6421
8.9992 | -101.6614
9.0031 | -101.692
9.0069 | -101.7366
9.0107 | -101.7934
9.0145 | -101.856
9.0183 | -101.9278
9.0222 | -102.0027
9.026 | -102.08
9.0298 | -102.1627
9.0336 | -102.2443
9.0374 | -102.3165
9.0413 | -102.384
9.0451 | -102.4488
9.0489 | -102.5019
9.0527 | -102.5473
9.0565 | -102.5868
9.0604 | -102.6252
9.0642 | -102.6609
9.068 | -102.7169
9.0718 | -102.7412
9.0756 | -102.7522
9.0794 | -102.7593
9.0833 | -102.7683
9.0871 | -102.7805
9.0909 | -102.7956
9.0947 | -102.804
9.0985 | -102.8113
9.1024 | -102.8221
9.1062 | -102.8349
9.11 | -102.845
9.1138 | -102.8501
9.1176 | -102.8569
9.1215 | -102.8659
9.1253 | -102.8753
9.1291 | -102.8834
9.1329 | -102.889
9.1367 | -102.8961
9.1406 | -102.9059
9.1444 | -102.9178
9.1482 | -102.9379
9.152 | -102.9568
9.1558 | -102.98
9.1597 | -103.001
9.1635 | -103.0189
9.1673 | -103.0423
9.1711 | -103.0649
9.1749 | -103.0845
9.1788 | -103.097
9.1826 | -103.1072
9.1864 | -103.119
9.1902 | -103.1294
9.194 | -103.1381
9.1979 | -103.1539
9.2017 | -103.168
9.2055 | -103.1812
9.2093 | -103.1857
9.2131 | -103.1957
9.217 | -103.2073
9.2208 | -103.2224
9.2246 | -103.2363
9.2284 | -103.2538
9.2322 | -103.2731
9.2361 | -103.2927
9.2399 | -103.3122
9.2437 | -103.3372
9.2475 | -103.3693
9.2513 | -103.4032
9.2552 | -103.4377
9.259 | -103.4754
9.2628 | -103.513
9.2666 | -103.5548
9.2704 | -103.5955
9.2743 | -103.632
9.2781 | -103.6627
9.2819 | -103.6844
9.2857 | -103.7064
9.2895 | -103.7235
9.2934 | -103.7343
9.2972 | -103.7437
9.301 | -103.7394
9.3048 | -103.7237
9.3086 | -103.6911
9.3125 | -103.6464
9.3163 | -103.5971
9.3201 | -103.5362
9.3239 | -103.4622
9.3277 | -103.3833
9.3316 | -103.2999
9.3354 | -103.2139
9.3392 | -103.1243
9.343 | -103.0395
9.3468 | -102.9575
9.3506 | -102.8755
9.3545 | -102.7988
9.3583 | -102.7305
9.3621 | -102.6722
9.3659 | -102.6724
9.3697 | -102.6474
9.3736 | -102.6359
9.3774 | -102.6383
9.3812 | -102.6547
9.385 | -102.6867
9.3888 | -102.7322
9.3927 | -102.7863
9.3965 | -102.8519
9.4003 | -102.9295
9.4041 | -103.0157
9.4079 | -103.1048
9.4118 | -103.1949
9.4156 | -103.2852
9.4194 | -103.3686
9.4232 | -103.4409
9.427 | -103.5051
9.4309 | -103.5622
9.4347 | -103.6015
9.4385 | -103.6296
9.4423 | -103.6564
9.4461 | -103.6766
9.45 | -103.6844
9.4538 | -103.6747
9.4576 | -103.656
9.4614 | -103.6356
9.4652 | -103.6107
9.4691 | -103.5801
9.4729 | -103.5473
9.4767 | -103.5149
9.4805 | -103.479
9.4843 | -103.4439
9.4882 | -103.4105
9.492 | -103.3775
9.4958 | -103.3511
9.4996 | -103.3216
9.5034 | -103.2945
9.5073 | -103.2684
9.5111 | -103.2445
9.5149 | -103.2239
9.5187 | -103.2073
9.5225 | -103.1918
9.5264 | -103.1751
9.5302 | -103.1588
9.534 | -103.1484
9.5378 | -103.1375
9.5416 | -103.1284
9.5455 | -103.1179
9.5493 | -103.1114
9.5531 | -103.1087
9.5569 | -103.1099
9.5607 | -103.1162
9.5646 | -103.1217
9.5684 | -103.128
9.5722 | -103.1337
9.576 | -103.1438
9.5798 | -103.1553
9.5837 | -103.1649
9.5875 | -103.1756
9.5913 | -103.1844
9.5951 | -103.1959
9.5989 | -103.2053
9.6028 | -103.2142
9.6066 | -103.2213
9.6104 | -103.2291
9.6142 | -103.2408
9.618 | -103.2531
9.6218 | -103.2643
9.6257 | -103.2725
9.6295 | -103.2834
9.6333 | -103.2957
9.6371 | -103.3045
9.6409 | -103.3127
9.6448 | -103.3133
9.6486 | -103.3111
9.6524 | -103.311
9.6562 | -103.3094
9.66 | -103.3033
9.6639 | -103.2959
9.6677 | -103.2854
9.6715 | -103.2787
9.6753 | -103.2701
9.6791 | -103.2589
9.683 | -103.2461
9.6868 | -103.2394
9.6906 | -103.2397
9.6944 | -103.2382
9.6982 | -103.234
9.7021 | -103.2331
9.7059 | -103.2401
9.7097 | -103.2517
9.7135 | -103.2675
9.7173 | -103.2823
9.7212 | -103.3006
9.725 | -103.3199
9.7288 | -103.3406
9.7326 | -103.3622
9.7364 | -103.3856
9.7403 | -103.4106
9.7441 | -103.4365
9.7479 | -103.4625
9.7517 | -103.4934
9.7555 | -103.5269
9.7594 | -103.5581
9.7632 | -103.5929
9.767 | -103.6332
9.7708 | -103.6737
9.7746 | -103.7181
9.7785 | -103.765
9.7823 | -103.8121
9.7861 | -103.8587
9.7899 | -103.8994
9.7937 | -103.9355
9.7976 | -103.9696
9.8014 | -104.0025
9.8052 | -104.0292
9.809 | -104.055
9.8128 | -104.0726
9.8167 | -104.0873
9.8205 | -104.1029
9.8243 | -104.1184
9.8281 | -104.1272
9.8319 | -104.1276
9.8358 | -104.1301
9.8396 | -104.1304
9.8434 | -104.1203
9.8472 | -104.1053
9.851 | -104.087
9.8549 | -104.0692
9.8587 | -104.0566
9.8625 | -104.0413
9.8663 | -104.0227
9.8701 | -104.0059
9.874 | -103.9953
9.8778 | -103.9844
9.8816 | -103.974
9.8854 | -103.9634
9.8892 | -103.9573
9.893 | -103.9479
9.8969 | -103.9386
9.9007 | -103.9364
9.9045 | -103.9354
9.9083 | -103.9328
9.9121 | -103.9255
9.916 | -103.9145
9.9198 | -103.9061
9.9236 | -103.8913
9.9274 | -103.8585
9.9312 | -103.8337
9.9351 | -103.8099
9.9389 | -103.7801
9.9427 | -103.7438
9.9465 | -103.7037
9.9503 | -103.6599
9.9542 | -103.614
9.958 | -103.5659
9.9618 | -103.5145
9.9656 | -103.4578
9.9694 | -103.4005
9.9733 | -103.3493
9.9771 | -103.3059
9.9809 | -103.2691
9.9847 | -103.2413
9.9885 | -103.2225
9.9924 | -103.2094
9.9962 | -103.2072
10 | -103.2163
10.0038 | -103.2366
10.0076 | -103.2697
10.0115 | -103.3132
10.0153 | -103.3666
10.0191 | -103.4256
10.0229 | -103.4915
10.0267 | -103.5647
10.0306 | -103.6456
10.0344 | -103.7268
10.0382 | -103.8085
10.042 | -103.8861
10.0458 | -103.9572
10.0497 | -104.0236
10.0535 | -104.0801
10.0573 | -104.1349
10.0611 | -104.1855
10.0649 | -104.222
10.0688 | -104.2486
10.0726 | -104.2725
10.0764 | -104.3042
10.0802 | -104.3106
10.084 | -104.3101
10.0879 | -104.3043
10.0917 | -104.2974
10.0955 | -104.289
10.0993 | -104.2832
10.1031 | -104.2781
10.107 | -104.2736
10.1108 | -104.2719
10.1146 | -104.2723
10.1184 | -104.2766
10.1222 | -104.2777
10.126 | -104.2782
10.1299 | -104.2764
10.1337 | -104.2731
10.1375 | -104.2735
10.1413 | -104.2757
10.1451 | -104.2703
10.149 | -104.2612
10.1528 | -104.2541
10.1566 | -104.2474
10.1604 | -104.2402
10.1642 | -104.2263
10.1681 | -104.2117
10.1719 | -104.1957
10.1757 | -104.175
10.1795 | -104.1461
10.1833 | -104.1126
10.1872 | -104.0826
10.191 | -104.0533
10.1948 | -104.0243
10.1986 | -103.9894
10.2024 | -103.9567
10.2063 | -103.9198
10.2101 | -103.8822
10.2139 | -103.8512
10.2177 | -103.8222
10.2215 | -103.7939
10.2254 | -103.7698
10.2292 | -103.7534
10.233 | -103.7402
10.2368 | -103.7317
10.2406 | -103.7283
10.2445 | -103.7312
10.2483 | -103.7388
10.2521 | -103.7517
10.2559 | -103.7694
10.2597 | -103.7927
10.2636 | -103.8234
10.2674 | -103.8569
10.2712 | -103.8898
10.275 | -103.923
10.2788 | -103.9635
10.2827 | -104.0063
10.2865 | -104.0435
10.2903 | -104.0789
10.2941 | -104.1156
10.2979 | -104.1493
10.3018 | -104.1712
10.3056 | -104.1878
10.3094 | -104.2043
10.3132 | -104.2216
10.317 | -104.2308
10.3209 | -104.233
10.3247 | -104.2328
10.3285 | -104.2363
10.3323 | -104.2382
10.3361 | -104.2321
10.34 | -104.2209
10.3438 | -104.2057
10.3476 | -104.1914
10.3514 | -104.1781
10.3552 | -104.1626
10.3591 | -104.1467
10.3629 | -104.1251
10.3667 | -104.1018
10.3705 | -104.0806
10.3743 | -104.0583
10.3782 | -104.0317
10.382 | -104.0024
10.3858 | -103.975
10.3896 | -103.9464
10.3934 | -103.9178
10.3972 | -103.8914
10.4011 | -103.8652
10.4049 | -103.8451
10.4087 | -103.8283
10.4125 | -103.8122
10.4163 | -103.7978
10.4202 | -103.7897
10.424 | -103.7857
10.4278 | -103.7819
10.4316 | -103.782
10.4354 | -103.7892
10.4393 | -103.7986
10.4431 | -103.8114
10.4469 | -103.8273
10.4507 | -103.8487
10.4545 | -103.8737
10.4584 | -103.897
10.4622 | -103.9198
10.466 | -103.9441
10.4698 | -103.9675
10.4736 | -103.9946
10.4775 | -104.0185
10.4813 | -104.0406
10.4851 | -104.0592
10.4889 | -104.0761
10.4927 | -104.0954
10.4966 | -104.113
10.5004 | -104.1279
10.5042 | -104.134
10.508 | -104.1364
10.5118 | -104.1363
10.5157 | -104.1387
10.5195 | -104.1382
10.5233 | -104.1363
10.5271 | -104.1289
10.5309 | -104.1263
10.5348 | -104.1267
10.5386 | -104.1236
10.5424 | -104.1205
10.5462 | -104.1171
10.55 | -104.114
10.5539 | -104.107
10.5577 | -104.0945
10.5615 | -104.0815
10.5653 | -104.0713
10.5691 | -104.0573
10.573 | -104.0449
10.5768 | -104.03
10.5806 | -104.0136
10.5844 | -103.9983
10.5882 | -103.9833
10.5921 | -103.9641
10.5959 | -103.9433
10.5997 | -103.9258
10.6035 | -103.9121
10.6073 | -103.9022
10.6112 | -103.8959
10.615 | -103.8882
10.6188 | -103.8839
10.6226 | -103.8832
10.6264 | -103.8854
10.6303 | -103.8911
10.6341 | -103.8993
10.6379 | -103.9089
10.6417 | -103.9287
10.6455 | -103.9504
10.6494 | -103.9719
10.6532 | -103.9979
10.657 | -104.0254
10.6608 | -104.057
10.6646 | -104.0917
10.6684 | -104.1235
10.6723 | -104.1539
10.6761 | -104.1837
10.6799 | -104.2119
10.6837 | -104.2414
10.6875 | -104.2641
10.6914 | -104.2848
10.6952 | -104.3065
10.699 | -104.3236
10.7028 | -104.335
10.7066 | -104.3391
10.7105 | -104.3458
10.7143 | -104.3543
10.7181 | -104.3563
10.7219 | -104.3528
10.7257 | -104.353
10.7296 | -104.3503
10.7334 | -104.3474
10.7372 | -104.3448
10.741 | -104.3405
10.7448 | -104.3351
10.7487 | -104.3335
10.7525 | -104.3349
10.7563 | -104.3334
10.7601 | -104.3258
10.7639 | -104.3145
10.7678 | -104.3086
10.7716 | -104.3068
10.7754 | -104.302
10.7792 | -104.2919
10.783 | -104.2812
10.7869 | -104.2757
10.7907 | -104.2731
10.7945 | -104.2679
10.7983 | -104.2601
10.8021 | -104.25
10.806 | -104.2394
10.8098 | -104.229
10.8136 | -104.2203
10.8174 | -104.2122
10.8212 | -104.2011
10.8251 | -104.1908
10.8289 | -104.1819
10.8327 | -104.1742
10.8365 | -104.1665
10.8403 | -104.1558
10.8442 | -104.1483
10.848 | -104.1421
10.8518 | -104.1346
10.8556 | -104.1298
10.8594 | -104.1282
10.8633 | -104.1304
10.8671 | -104.1317
10.8709 | -104.1308
10.8747 | -104.1309
10.8785 | -104.1337
10.8824 | -104.137
10.8862 | -104.1386
10.89 | -104.144
10.8938 | -104.1496
10.8976 | -104.1551
10.9015 | -104.1615
10.9053 | -104.1662
10.9091 | -104.1667
10.9129 | -104.1609
10.9167 | -104.1556
10.9206 | -104.1459
10.9244 | -104.1238
10.9282 | -104.1
10.932 | -104.0773
10.9358 | -104.055
10.9396 | -104.0315
10.9435 | -104.0052
10.9473 | -103.9769
10.9511 | -103.9509
10.9549 | -103.9243
10.9587 | -103.9028
10.9626 | -103.8843
10.9664 | -103.8641
10.9702 | -103.8471
10.974 | -103.8344
10.9778 | -103.8245
10.9817 | -103.8133
10.9855 | -103.8046
10.9893 | -103.8012
10.9931 | -103.8014
10.9969 | -103.8076
11.0008 | -103.8186
11.0046 | -103.8343
11.0084 | -103.8549
11.0122 | -103.8793
11.016 | -103.906
11.0199 | -103.937
11.0237 | -103.9622
11.0275 | -103.9848
11.0313 | -104.009
11.0351 | -104.0322
11.039 | -104.053
11.0428 | -104.0699
11.0466 | -104.0879
11.0504 | -104.1019
11.0542 | -104.1104
11.0581 | -104.1153
11.0619 | -104.1157
11.0657 | -104.1153
11.0695 | -104.1137
11.0733 | -104.1112
11.0772 | -104.1084
11.081 | -104.106
11.0848 | -104.1047
11.0886 | -104.1013
11.0924 | -104.0983
11.0963 | -104.0967
11.1001 | -104.0962
11.1039 | -104.0951
11.1077 | -104.1008
11.1115 | -104.112
11.1154 | -104.125
11.1192 | -104.1393
11.123 | -104.1554
11.1268 | -104.1767
11.1306 | -104.1968
11.1345 | -104.2136
11.1383 | -104.2357
11.1421 | -104.2624
11.1459 | -104.2867
11.1497 | -104.304
11.1536 | -104.3182
11.1574 | -104.3301
11.1612 | -104.339
11.165 | -104.3439
11.1688 | -104.3483
11.1727 | -104.351
11.1765 | -104.3507
11.1803 | -104.3478
11.1841 | -104.3453
11.1879 | -104.3431
11.1917 | -104.3422
11.1956 | -104.3331
11.1994 | -104.3227
11.2032 | -104.3119
11.207 | -104.2995
11.2108 | -104.2869
11.2147 | -104.2736
11.2185 | -104.2646
11.2223 | -104.2555
11.2261 | -104.2475
11.2299 | -104.2378
11.2338 | -104.23
11.2376 | -104.2259
11.2414 | -104.2135
11.2452 | -104.2038
11.249 | -104.1982
11.2529 | -104.1919
11.2567 | -104.1878
11.2605 | -104.1837
11.2643 | -104.1812
11.2681 | -104.1778
11.272 | -104.1705
11.2758 | -104.1634
11.2796 | -104.1611
11.2834 | -104.1598
11.2872 | -104.1563
11.2911 | -104.1523
11.2949 | -104.1513
11.2987 | -104.152
11.3025 | -104.1492
11.3063 | -104.1463
11.3102 | -104.1444
11.314 | -104.1449
11.3178 | -104.1475
11.3216 | -104.1484
11.3254 | -104.1532
11.3293 | -104.1617
11.3331 | -104.1735
11.3369 | -104.1885
11.3407 | -104.2099
11.3445 | -104.233
11.3484 | -104.2575
11.3522 | -104.2831
11.356 | -104.3126
11.3598 | -104.3479
11.3636 | -104.3842
11.3675 | -104.4184
11.3713 | -104.4511
11.3751 | -104.4802
11.3789 | -104.5038
11.3827 | -104.5246
11.3866 | -104.5409
11.3904 | -104.5461
11.3942 | -104.5464
11.398 | -104.5412
11.4018 | -104.5276
11.4057 | -104.5177
11.4095 | -104.4948
11.4133 | -104.4634
11.4171 | -104.4289
11.4209 | -104.3897
11.4248 | -104.3477
11.4286 | -104.3053
11.4324 | -104.2605
11.4362 | -104.2142
11.44 | -104.1663
11.4438 | -104.1165
11.4477 | -104.0667
11.4515 | -104.0232
11.4553 | -103.982
11.4591 | -103.9426
11.4629 | -103.9088
11.4668 | -103.8775
11.4706 | -103.8517
11.4744 | -103.8284
11.4782 | -103.8088
11.482 | -103.7923
11.4859 | -103.7803
11.4897 | -103.7661
11.4935 | -103.7547
11.4973 | -103.7482
11.5011 | -103.7473
11.505 | -103.7485
11.5088 | -103.7545
11.5126 | -103.7602
11.5164 | -103.7659
11.5202 | -103.7772
11.5241 | -103.7917
11.5279 | -103.8063
11.5317 | -103.8231
11.5355 | -103.8403
11.5393 | -103.8616
11.5432 | -103.8847
11.547 | -103.9048
11.5508 | -103.924
11.5546 | -103.9501
11.5584 | -103.9764
11.5623 | -103.9955
11.5661 | -104.0154
11.5699 | -104.0334
11.5737 | -104.046
11.5775 | -104.0631
11.5814 | -104.0797
11.5852 | -104.0905
11.589 | -104.0988
11.5928 | -104.109
11.5966 | -104.12
11.6005 | -104.1326
11.6043 | -104.1442
11.6081 | -104.1545
11.6119 | -104.166
11.6157 | -104.1793
11.6196 | -104.1897
11.6234 | -104.2011
11.6272 | -104.2113
11.631 | -104.221
11.6348 | -104.2325
11.6387 | -104.2414
11.6425 | -104.2512
11.6463 | -104.2625
11.6501 | -104.2736
11.6539 | -104.2809
11.6578 | -104.2865
11.6616 | -104.2943
11.6654 | -104.2954
11.6692 | -104.2917
11.673 | -104.2931
11.6769 | -104.2959
11.6807 | -104.2987
11.6845 | -104.3027
11.6883 | -104.3021
11.6921 | -104.3021
11.696 | -104.3047
11.6998 | -104.3083
11.7036 | -104.3161
11.7074 | -104.3213
11.7112 | -104.3269
11.715 | -104.3323
11.7189 | -104.3403
11.7227 | -104.3421
11.7265 | -104.3408
11.7303 | -104.3458
11.7341 | -104.3498
11.738 | -104.3513
11.7418 | -104.3542
11.7456 | -104.3581
11.7494 | -104.3577
11.7532 | -104.3581
11.7571 | -104.3641
11.7609 | -104.3671
11.7647 | -104.3654
11.7685 | -104.361
11.7723 | -104.3602
11.7762 | -104.3618
11.78 | -104.3624
11.7838 | -104.3646
11.7876 | -104.368
11.7914 | -104.3724
11.7953 | -104.3781
11.7991 | -104.3878
11.8029 | -104.3963
11.8067 | -104.407
11.8105 | -104.4171
11.8144 | -104.4271
11.8182 | -104.4371
11.822 | -104.4481
11.8258 | -104.4643
11.8296 | -104.4819
11.8335 | -104.4973
11.8373 | -104.5171
11.8411 | -104.5376
11.8449 | -104.5522
11.8487 | -104.5673
11.8526 | -104.5817
11.8564 | -104.595
11.8602 | -104.6069
11.864 | -104.6146
11.8678 | -104.622
11.8717 | -104.6261
11.8755 | -104.6294
11.8793 | -104.6295
11.8831 | -104.6284
11.8869 | -104.6261
11.8908 | -104.6186
11.8946 | -104.6114
11.8984 | -104.6038
11.9022 | -104.5905
11.906 | -104.5755
11.9099 | -104.5588
11.9137 | -104.5399
11.9175 | -104.5177
11.9213 | -104.5011
11.9251 | -104.4828
11.929 | -104.4626
11.9328 | -104.4423
11.9366 | -104.4281
11.9404 | -104.4161
11.9442 | -104.4044
11.9481 | -104.3927
11.9519 | -104.382
11.9557 | -104.3717
11.9595 | -104.3676
11.9633 | -104.3679
11.9672 | -104.3699
11.971 | -104.3757
11.9748 | -104.3817
11.9786 | -104.3894
11.9824 | -104.4009
11.9862 | -104.4135
11.9901 | -104.4294
11.9939 | -104.4474
11.9977 | -104.4657
12.0015 | -104.4843
12.0053 | -104.502
12.0092 | -104.5213
12.013 | -104.5406
12.0168 | -104.5584
12.0206 | -104.5718
12.0244 | -104.5878
12.0283 | -104.6001
12.0321 | -104.6114
12.0359 | -104.6222
12.0397 | -104.6288
12.0435 | -104.6335
12.0474 | -104.6336
12.0512 | -104.6355
12.055 | -104.637
12.0588 | -104.6376
12.0626 | -104.6423
12.0665 | -104.6468
12.0703 | -104.6509
12.0741 | -104.6596
12.0779 | -104.6706
12.0817 | -104.6844
12.0856 | -104.7027
12.0894 | -104.7202
12.0932 | -104.7336
12.097 | -104.7493
12.1008 | -104.7636
12.1047 | -104.7799
12.1085 | -104.7945
12.1123 | -104.8054
12.1161 | -104.8151
12.1199 | -104.8249
12.1238 | -104.8381
12.1276 | -104.8514
12.1314 | -104.8642
12.1352 | -104.8698
12.139 | -104.8739
12.1429 | -104.8814
12.1467 | -104.8874
12.1505 | -104.8954
12.1543 | -104.9007
12.1581 | -104.905
12.162 | -104.9105
12.1658 | -104.9137
12.1696 | -104.9134
12.1734 | -104.9131
12.1772 | -104.9143
12.1811 | -104.9155
12.1849 | -104.9131
12.1887 | -104.9118
12.1925 | -104.9135
12.1963 | -104.9143
12.2002 | -104.9114
12.204 | -104.9057
12.2078 | -104.8996
12.2116 | -104.894
12.2154 | -104.8865
12.2193 | -104.8804
12.2231 | -104.8773
12.2269 | -104.8744
12.2307 | -104.8701
12.2345 | -104.8659
12.2384 | -104.8605
12.2422 | -104.8568
12.246 | -104.8525
12.2498 | -104.8512
12.2536 | -104.851
12.2574 | -104.8503
12.2613 | -104.8483
12.2651 | -104.8492
12.2689 | -104.849
12.2727 | -104.8496
12.2765 | -104.8508
12.2804 | -104.8503
12.2842 | -104.8474
12.288 | -104.8449
12.2918 | -104.8424
12.2956 | -104.8374
12.2995 | -104.8329
12.3033 | -104.8284
12.3071 | -104.8247
12.3109 | -104.8218
12.3147 | -104.8203
12.3186 | -104.8149
12.3224 | -104.8088
12.3262 | -104.8025
12.33 | -104.7965
12.3338 | -104.7913
12.3377 | -104.7814
12.3415 | -104.7682
12.3453 | -104.758
12.3491 | -104.7495
12.3529 | -104.7394
12.3568 | -104.7284
12.3606 | -104.7144
12.3644 | -104.6996
12.3682 | -104.6868
12.372 | -104.6747
12.3759 | -104.6659
12.3797 | -104.6557
12.3835 | -104.6462
12.3873 | -104.6393
12.3911 | -104.6363
12.395 | -104.6338
12.3988 | -104.6333
12.4026 | -104.6313
12.4064 | -104.6322
12.4102 | -104.6361
12.4141 | -104.6371
12.4179 | -104.6397
12.4217 | -104.6431
12.4255 | -104.6469
12.4293 | -104.651
12.4332 | -104.6557
12.437 | -104.6609
12.4408 | -104.6697
12.4446 | -104.6773
12.4484 | -104.6826
12.4523 | -104.6882
12.4561 | -104.6948
12.4599 | -104.7021
12.4637 | -104.7087
12.4675 | -104.7147
12.4714 | -104.7199
12.4752 | -104.7256
12.479 | -104.729
12.4828 | -104.7355
12.4866 | -104.7424
12.4905 | -104.7492
12.4943 | -104.7588
12.4981 | -104.7701
12.5019 | -104.7784
12.5057 | -104.7873
12.5095 | -104.7955
12.5134 | -104.8016
12.5172 | -104.811
12.521 | -104.8206
12.5248 | -104.8257
12.5286 | -104.8298
12.5325 | -104.8344
12.5363 | -104.838
12.5401 | -104.8415
12.5439 | -104.8456
12.5477 | -104.8443
12.5516 | -104.8399
12.5554 | -104.8331
12.5592 | -104.8277
12.563 | -104.824
12.5668 | -104.8151
12.5707 | -104.8025
12.5745 | -104.794
12.5783 | -104.7841
12.5821 | -104.7744
12.5859 | -104.7647
12.5898 | -104.7555
12.5936 | -104.75
12.5974 | -104.7435
12.6012 | -104.7363
12.605 | -104.731
12.6089 | -104.7256
12.6127 | -104.7213
12.6165 | -104.7203
12.6203 | -104.7177
12.6241 | -104.7179
12.628 | -104.7215
12.6318 | -104.7262
12.6356 | -104.7325
12.6394 | -104.7402
12.6432 | -104.7485
12.6471 | -104.7584
12.6509 | -104.7657
12.6547 | -104.773
12.6585 | -104.781
12.6623 | -104.7879
12.6662 | -104.7986
12.67 | -104.8068
12.6738 | -104.8098
12.6776 | -104.8179
12.6814 | -104.8255
12.6853 | -104.8301
12.6891 | -104.831
12.6929 | -104.8286
12.6967 | -104.8284
12.7005 | -104.8268
12.7044 | -104.8217
12.7082 | -104.8123
12.712 | -104.801
12.7158 | -104.7863
12.7196 | -104.7702
12.7235 | -104.7517
12.7273 | -104.7302
12.7311 | -104.7093
12.7349 | -104.6896
12.7387 | -104.6698
12.7426 | -104.6519
12.7464 | -104.6328
12.7502 | -104.6127
12.754 | -104.5949
12.7578 | -104.5788
12.7616 | -104.566
12.7655 | -104.5608
12.7693 | -104.5565
12.7731 | -104.5559
12.7769 | -104.5579
12.7807 | -104.5589
12.7846 | -104.5617
12.7884 | -104.5645
12.7922 | -104.5684
12.796 | -104.5751
12.7998 | -104.5791
12.8037 | -104.5861
12.8075 | -104.5928
12.8113 | -104.5986
12.8151 | -104.6028
12.8189 | -104.605
12.8228 | -104.6047
12.8266 | -104.6029
12.8304 | -104.6008
12.8342 | -104.5983
12.838 | -104.5934
12.8419 | -104.5864
12.8457 | -104.5777
12.8495 | -104.5666
12.8533 | -104.5567
12.8571 | -104.5472
12.861 | -104.5332
12.8648 | -104.5188
12.8686 | -104.508
12.8724 | -104.499
12.8762 | -104.4874
12.8801 | -104.4779
12.8839 | -104.4652
12.8877 | -104.4536
12.8915 | -104.4458
12.8953 | -104.4402
12.8992 | -104.4366
12.903 | -104.4346
12.9068 | -104.4264
12.9106 | -104.4218
12.9144 | -104.4215
12.9183 | -104.4206
12.9221 | -104.4202
12.9259 | -104.4201
12.9297 | -104.4195
12.9335 | -104.4223
12.9374 | -104.4297
12.9412 | -104.4368
12.945 | -104.4452
12.9488 | -104.455
12.9526 | -104.4613
12.9565 | -104.4701
12.9603 | -104.4805
12.9641 | -104.4938
12.9679 | -104.5098
12.9717 | -104.5241
12.9756 | -104.5418
12.9794 | -104.5572
12.9832 | -104.5679
12.987 | -104.581
12.9908 | -104.5937
12.9947 | -104.6063
12.9985 | -104.6168
13.0023 | -104.6258
13.0061 | -104.6347
13.0099 | -104.6442
13.0138 | -104.6542
13.0176 | -104.6579
13.0214 | -104.6632
13.0252 | -104.6654
13.029 | -104.6668
13.0328 | -104.6668
13.0367 | -104.6634
13.0405 | -104.6609
13.0443 | -104.6536
13.0481 | -104.6514
13.0519 | -104.6476
13.0558 | -104.6386
13.0596 | -104.632
13.0634 | -104.6274
13.0672 | -104.6214
13.071 | -104.6144
13.0749 | -104.6069
13.0787 | -104.5982
13.0825 | -104.5869
13.0863 | -104.5718
13.0901 | -104.5582
13.094 | -104.5487
13.0978 | -104.5353
13.1016 | -104.5179
13.1054 | -104.4995
13.1092 | -104.4759
13.1131 | -104.4544
13.1169 | -104.4313
13.1207 | -104.4047
13.1245 | -104.3737
13.1283 | -104.342
13.1322 | -104.3092
13.136 | -104.2761
13.1398 | -104.2424
13.1436 | -104.2068
13.1474 | -104.1732
13.1513 | -104.1412
13.1551 | -104.1098
13.1589 | -104.0818
13.1627 | -104.0534
13.1665 | -104.0262
13.1704 | -104.0048
13.1742 | -103.9858
13.178 | -103.9687
13.1818 | -103.9567
13.1856 | -103.9477
13.1895 | -103.9419
13.1933 | -103.9401
13.1971 | -103.9412
13.2009 | -103.9456
13.2047 | -103.9577
13.2086 | -103.9688
13.2124 | -103.9825
13.2162 | -103.9977
13.22 | -104.0159
13.2238 | -104.0351
13.2277 | -104.0541
13.2315 | -104.0722
13.2353 | -104.0937
13.2391 | -104.1138
13.2429 | -104.1323
13.2468 | -104.154
13.2506 | -104.1706
13.2544 | -104.1873
13.2582 | -104.2034
13.262 | -104.2172
13.2659 | -104.2268
13.2697 | -104.2386
13.2735 | -104.2546
13.2773 | -104.2686
13.2811 | -104.2765
13.285 | -104.2853
13.2888 | -104.2904
13.2926 | -104.2935
13.2964 | -104.2978
13.3002 | -104.3034
13.304 | -104.3084
13.3079 | -104.3145
13.3117 | -104.318
13.3155 | -104.3253
13.3193 | -104.3318
13.3231 | -104.3381
13.327 | -104.3429
13.3308 | -104.3506
13.3346 | -104.3607
13.3384 | -104.3673
13.3422 | -104.3722
13.3461 | -104.3776
13.3499 | -104.3822
13.3537 | -104.3879
13.3575 | -104.396
13.3613 | -104.4013
13.3652 | -104.4046
13.369 | -104.4099
13.3728 | -104.4131
13.3766 | -104.4156
13.3804 | -104.4181
13.3843 | -104.419
13.3881 | -104.4193
13.3919 | -104.4179
13.3957 | -104.4169
13.3995 | -104.4163
13.4034 | -104.4157
13.4072 | -104.4149
13.411 | -104.4112
13.4148 | -104.4121
13.4186 | -104.4159
13.4225 | -104.4178
13.4263 | -104.4199
13.4301 | -104.4209
13.4339 | -104.4207
13.4377 | -104.4232
13.4416 | -104.4278
13.4454 | -104.4368
13.4492 | -104.445
13.453 | -104.4533
13.4568 | -104.4642
13.4607 | -104.476
13.4645 | -104.4852
13.4683 | -104.4942
13.4721 | -104.5072
13.4759 | -104.5207
13.4798 | -104.5326
13.4836 | -104.545
13.4874 | -104.5573
13.4912 | -104.5693
13.495 | -104.5847
13.4989 | -104.5973
13.5027 | -104.6116
13.5065 | -104.6285
13.5103 | -104.6423
13.5141 | -104.6574
13.518 | -104.6727
13.5218 | -104.6905
13.5256 | -104.7033
13.5294 | -104.7184
13.5332 | -104.7352
13.5371 | -104.75
13.5409 | -104.7686
13.5447 | -104.7906
13.5485 | -104.8084
13.5523 | -104.8243
13.5562 | -104.8417
13.56 | -104.8587
13.5638 | -104.8729
13.5676 | -104.8894
13.5714 | -104.9094
13.5752 | -104.9297
13.5791 | -104.9454
13.5829 | -104.9577
13.5867 | -104.9683
13.5905 | -104.9785
13.5943 | -104.9902
13.5982 | -105.0018
13.602 | -105.01
13.6058 | -105.0184
13.6096 | -105.0183
13.6134 | -105.0166
13.6173 | -105.0123
13.6211 | -105.0082
13.6249 | -105.0042
13.6287 | -104.9996
13.6325 | -104.9935
13.6364 | -104.9799
13.6402 | -104.961
13.644 | -104.9376
13.6478 | -104.9101
13.6516 | -104.8712
13.6555 | -104.8078
13.6593 | -104.7442
13.6631 | -104.6793
13.6669 | -104.6059
13.6707 | -104.5334
13.6746 | -104.456
13.6784 | -104.3752
13.6822 | -104.2867
13.686 | -104.1893
13.6898 | -104.0854
13.6937 | -103.976
13.6975 | -103.8583
13.7013 | -103.7383
13.7051 | -103.6175
13.7089 | -103.4982
13.7128 | -103.3803
13.7166 | -103.2645
13.7204 | -103.1569
13.7242 | -103.0545
13.728 | -102.9585
13.7319 | -102.8749
13.7357 | -102.8008
13.7395 | -102.7375
13.7433 | -102.6876
13.7471 | -102.6518
13.751 | -102.6297
13.7548 | -102.6218
13.7586 | -102.6291
13.7624 | -102.6502
13.7662 | -102.6846
13.7701 | -102.7312
13.7739 | -102.8383
13.7777 | -102.9055
13.7815 | -102.9816
13.7853 | -103.0663
13.7892 | -103.1555
13.793 | -103.2503
13.7968 | -103.3512
13.8006 | -103.4553
13.8044 | -103.56
13.8083 | -103.6667
13.8121 | -103.7698
13.8159 | -103.8689
13.8197 | -103.9637
13.8235 | -104.0539
13.8273 | -104.1381
13.8312 | -104.215
13.835 | -104.2865
13.8388 | -104.3507
13.8426 | -104.4082
13.8464 | -104.4576
13.8503 | -104.5057
13.8541 | -104.5594
13.8579 | -104.5947
13.8617 | -104.6157
13.8655 | -104.6307
13.8694 | -104.6363
13.8732 | -104.6369
13.877 | -104.6304
13.8808 | -104.6235
13.8846 | -104.6188
13.8885 | -104.6121
13.8923 | -104.6072
13.8961 | -104.6019
13.8999 | -104.5862
13.9037 | -104.5751
13.9076 | -104.5666
13.9114 | -104.5581
13.9152 | -104.5491
13.919 | -104.5382
13.9228 | -104.5289
13.9267 | -104.5217
13.9305 | -104.5123
13.9343 | -104.5042
13.9381 | -104.4957
13.9419 | -104.487
13.9458 | -104.4772
13.9496 | -104.4703
13.9534 | -104.4626
13.9572 | -104.4503
13.961 | -104.4357
13.9649 | -104.4228
13.9687 | -104.4122
13.9725 | -104.4024
13.9763 | -104.391
13.9801 | -104.3787
13.984 | -104.3703
13.9878 | -104.3614
13.9916 | -104.3534
13.9954 | -104.3481
13.9992 | -104.3407
14.0031 | -104.3368
14.0069 | -104.3378
14.0107 | -104.3394
14.0145 | -104.3459
14.0183 | -104.3565
14.0222 | -104.37
14.026 | -104.3848
14.0298 | -104.4009
14.0336 | -104.4201
14.0374 | -104.4439
14.0413 | -104.4701
14.0451 | -104.496
14.0489 | -104.5219
14.0527 | -104.5456
14.0565 | -104.5709
14.0604 | -104.5983
14.0642 | -104.6241
14.068 | -104.6509
14.0718 | -104.6766
14.0756 | -104.7017
14.0794 | -104.7281
14.0833 | -104.7495
14.0871 | -104.7679
14.0909 | -104.7861
14.0947 | -104.8019
14.0985 | -104.8206
14.1024 | -104.8477
14.1062 | -104.8658
14.11 | -104.8774
14.1138 | -104.8885
14.1176 | -104.9007
14.1215 | -104.9123
14.1253 | -104.9225
14.1291 | -104.9351
14.1329 | -104.9458
14.1367 | -104.9548
14.1406 | -104.9681
14.1444 | -104.9835
14.1482 | -104.9996
14.152 | -105.0161
14.1558 | -105.0311
14.1597 | -105.048
14.1635 | -105.0679
14.1673 | -105.0842
14.1711 | -105.099
14.1749 | -105.1142
14.1788 | -105.129
14.1826 | -105.1452
14.1864 | -105.1581
14.1902 | -105.1692
14.194 | -105.181
14.1979 | -105.1915
14.2017 | -105.2031
14.2055 | -105.213
14.2093 | -105.2168
14.2131 | -105.2231
14.217 | -105.2267
14.2208 | -105.2275
14.2246 | -105.2251
14.2284 | -105.225
14.2322 | -105.2285
14.2361 | -105.2322
14.2399 | -105.2359
14.2437 | -105.2397
14.2475 | -105.2464
14.2513 | -105.2529
14.2552 | -105.2592
14.259 | -105.2663
14.2628 | -105.2725
14.2666 | -105.2797
14.2704 | -105.2866
14.2743 | -105.2953
14.2781 | -105.3025
14.2819 | -105.3098
14.2857 | -105.3218
14.2895 | -105.3338
14.2934 | -105.3439
14.2972 | -105.354
14.301 | -105.3592
14.3048 | -105.3612
14.3086 | -105.3635
14.3125 | -105.3645
14.3163 | -105.3644
14.3201 | -105.3653
14.3239 | -105.3596
14.3277 | -105.3597
14.3316 | -105.3592
14.3354 | -105.3558
14.3392 | -105.3554
14.343 | -105.3529
14.3468 | -105.3465
14.3506 | -105.338
14.3545 | -105.3296
14.3583 | -105.3206
14.3621 | -105.3149
14.3659 | -105.3062
14.3697 | -105.2962
14.3736 | -105.2829
14.3774 | -105.2693
14.3812 | -105.2553
14.385 | -105.2403
14.3888 | -105.2267
14.3927 | -105.213
14.3965 | -105.1996
14.4003 | -105.1854
14.4041 | -105.1687
14.4079 | -105.1507
14.4118 | -105.1325
14.4156 | -105.12
14.4194 | -105.1055
14.4232 | -105.0917
14.427 | -105.0793
14.4309 | -105.0722
14.4347 | -105.062
14.4385 | -105.0498
14.4423 | -105.0414
14.4461 | -105.0348
14.45 | -105.0321
14.4538 | -105.0299
14.4576 | -105.0273
14.4614 | -105.0287
14.4652 | -105.0314
14.4691 | -105.0358
14.4729 | -105.0424
14.4767 | -105.0494
14.4805 | -105.0561
14.4843 | -105.0666
14.4882 | -105.0776
14.492 | -105.0856
14.4958 | -105.0923
14.4996 | -105.1004
14.5034 | -105.1113
14.5073 | -105.1238
14.5111 | -105.1346
14.5149 | -105.1438
14.5187 | -105.1529
14.5225 | -105.164
14.5264 | -105.1731
14.5302 | -105.1782
14.534 | -105.1833
14.5378 | -105.1844
14.5416 | -105.1832
14.5455 | -105.1828
14.5493 | -105.1817
14.5531 | -105.1797
14.5569 | -105.1754
14.5607 | -105.1698
14.5646 | -105.1639
14.5684 | -105.1551
14.5722 | -105.1441
14.576 | -105.1382
14.5798 | -105.1293
14.5837 | -105.1223
14.5875 | -105.1154
14.5913 | -105.1065
14.5951 | -105.0995
14.5989 | -105.0922
14.6028 | -105.0854
14.6066 | -105.0805
14.6104 | -105.0798
14.6142 | -105.0824
14.618 | -105.0859
14.6218 | -105.091
14.6257 | -105.0927
14.6295 | -105.1002
14.6333 | -105.1128
14.6371 | -105.1259
14.6409 | -105.1378
14.6448 | -105.1556
14.6486 | -105.1755
14.6524 | -105.1937
14.6562 | -105.215
14.66 | -105.2365
14.6639 | -105.2585
14.6677 | -105.2826
14.6715 | -105.3073
14.6753 | -105.3329
14.6791 | -105.3572
14.683 | -105.378
14.6868 | -105.3963
14.6906 | -105.4157
14.6944 | -105.4329
14.6982 | -105.4452
14.7021 | -105.4555
14.7059 | -105.4641
14.7097 | -105.4703
14.7135 | -105.4703
14.7173 | -105.4685
14.7212 | -105.4637
14.725 | -105.4573
14.7288 | -105.4477
14.7326 | -105.4331
14.7364 | -105.4142
14.7403 | -105.3972
14.7441 | -105.3771
14.7479 | -105.3582
14.7517 | -105.3382
14.7555 | -105.313
14.7594 | -105.2841
14.7632 | -105.2602
14.767 | -105.2386
14.7708 | -105.2181
14.7746 | -105.1927
14.7785 | -105.1678
14.7823 | -105.1482
14.7861 | -105.1277
14.7899 | -105.1085
14.7937 | -105.0919
14.7976 | -105.0774
14.8014 | -105.0656
14.8052 | -105.0563
14.809 | -105.0501
14.8128 | -105.0443
14.8167 | -105.0385
14.8205 | -105.0331
14.8243 | -105.0314
14.8281 | -105.0341
14.8319 | -105.0391
14.8358 | -105.0477
14.8396 | -105.0525
14.8434 | -105.0573
14.8472 | -105.0631
14.851 | -105.0693
14.8549 | -105.0773
14.8587 | -105.0876
14.8625 | -105.101
14.8663 | -105.1105
14.8701 | -105.1209
14.874 | -105.1339
14.8778 | -105.1449
14.8816 | -105.1565
14.8854 | -105.1673
14.8892 | -105.1736
14.893 | -105.1774
14.8969 | -105.1843
14.9007 | -105.1946
14.9045 | -105.2018
14.9083 | -105.2079
14.9121 | -105.2156
14.916 | -105.222
14.9198 | -105.2229
14.9236 | -105.2232
14.9274 | -105.2226
14.9312 | -105.2179
14.9351 | -105.2125
14.9389 | -105.2036
14.9427 | -105.1938
14.9465 | -105.1843
14.9503 | -105.1756
14.9542 | -105.1649
14.958 | -105.1505
14.9618 | -105.133
14.9656 | -105.1171
14.9694 | -105.1033
14.9733 | -105.0903
14.9771 | -105.0732
14.9809 | -105.057
14.9847 | -105.0426
14.9885 | -105.03
14.9924 | -105.0175
14.9962 | -105.0032
15 | -104.9913
15.0038 | -104.9813
15.0076 | -104.9738
15.0115 | -104.9675
15.0153 | -104.963
15.0191 | -104.9586
15.0229 | -104.954
15.0267 | -104.9556
15.0306 | -104.9613
15.0344 | -104.9661
15.0382 | -104.9698
15.042 | -104.9736
15.0458 | -104.9782
15.0497 | -104.9853
15.0535 | -104.9944
15.0573 | -105.0032
15.0611 | -105.013
15.0649 | -105.0231
15.0688 | -105.0307
15.0726 | -105.0413
15.0764 | -105.0502
15.0802 | -105.0609
15.084 | -105.07
15.0879 | -105.078
15.0917 | -105.0857
15.0955 | -105.0903
15.0993 | -105.095
15.1031 | -105.0982
15.107 | -105.1026
15.1108 | -105.1089
15.1146 | -105.1092
15.1184 | -105.1038
15.1222 | -105.0955
15.126 | -105.0862
15.1299 | -105.0775
15.1337 | -105.0657
15.1375 | -105.0541
15.1413 | -105.0395
15.1451 | -105.0241
15.149 | -105.0099
15.1528 | -104.996
15.1566 | -104.9803
15.1604 | -104.9635
15.1642 | -104.9443
15.1681 | -104.9256
15.1719 | -104.904
15.1757 | -104.8829
15.1795 | -104.8642
15.1833 | -104.8428
15.1872 | -104.8208
15.191 | -104.8004
15.1948 | -104.7789
15.1986 | -104.7599
15.2024 | -104.7423
15.2063 | -104.7267
15.2101 | -104.7124
15.2139 | -104.7
15.2177 | -104.6883
15.2215 | -104.6795
15.2254 | -104.6717
15.2292 | -104.6632
15.233 | -104.6533
15.2368 | -104.6459
15.2406 | -104.6413
15.2445 | -104.6359
15.2483 | -104.6304
15.2521 | -104.6274
15.2559 | -104.6271
15.2597 | -104.6298
15.2636 | -104.6326
15.2674 | -104.6337
15.2712 | -104.6392
15.275 | -104.6456
15.2788 | -104.651
15.2827 | -104.6576
15.2865 | -104.6632
15.2903 | -104.6658
15.2941 | -104.6712
15.2979 | -104.6765
15.3018 | -104.682
15.3056 | -104.6888
15.3094 | -104.6937
15.3132 | -104.7006
15.317 | -104.7043
15.3209 | -104.7053
15.3247 | -104.7046
15.3285 | -104.705
15.3323 | -104.7079
15.3361 | -104.7086
15.34 | -104.7093
15.3438 | -104.7053
15.3476 | -104.7005
15.3514 | -104.6947
15.3552 | -104.688
15.3591 | -104.6795
15.3629 | -104.6685
15.3667 | -104.6584
15.3705 | -104.6479
15.3743 | -104.6386
15.3782 | -104.629
15.382 | -104.6185
15.3858 | -104.6061
15.3896 | -104.5962
15.3934 | -104.5864
15.3972 | -104.574
15.4011 | -104.562
15.4049 | -104.551
15.4087 | -104.5402
15.4125 | -104.5319
15.4163 | -104.5218
15.4202 | -104.5098
15.424 | -104.4979
15.4278 | -104.4933
15.4316 | -104.4855
15.4354 | -104.4781
15.4393 | -104.4725
15.4431 | -104.4674
15.4469 | -104.4617
15.4507 | -104.461
15.4545 | -104.4626
15.4584 | -104.4619
15.4622 | -104.462
15.466 | -104.4629
15.4698 | -104.4661
15.4736 | -104.4689
15.4775 | -104.4725
15.4813 | -104.4764
15.4851 | -104.4824
15.4889 | -104.4907
15.4927 | -104.4962
15.4966 | -104.5022
15.5004 | -104.513
15.5042 | -104.5215
15.508 | -104.5173
15.5118 | -103.9503
15.5157 | -103.1657
15.5195 | -102.8908
15.5233 | -102.6326
15.5271 | -102.3532
15.5309 | -102.0566
15.5348 | -101.7438
15.5386 | -101.4177
15.5424 | -101.0809
15.5462 | -100.7378
15.55 | -100.3911
15.5539 | -100.0439
15.5577 | -99.7007
15.5615 | -99.3641
15.5653 | -99.0368
15.5691 | -98.7199
15.573 | -98.4167
15.5768 | -98.1284
15.5806 | -97.8573
15.5844 | -97.6044
15.5882 | -97.3707
15.5921 | -97.1575
15.5959 | -96.9658
15.5997 | -96.7961
15.6035 | -96.6497
15.6073 | -96.5268
15.6112 | -96.4275
15.615 | -96.3522
15.6188 | -96.3009
15.6226 | -96.2738
15.6264 | -96.2715
15.6303 | -96.2935
15.6341 | -96.3397
15.6379 | -96.4106
15.6417 | -96.5061
15.6455 | -96.6259
15.6494 | -96.7688
15.6532 | -96.9345
15.657 | -97.1231
15.6608 | -97.3335
15.6646 | -97.5647
15.6684 | -97.8163
15.6723 | -98.0873
15.6761 | -98.3757
15.6799 | -98.6804
15.6837 | -99.0009
15.6875 | -99.3343
15.6914 | -99.6782
15.6952 | -100.0302
15.699 | -100.3881
15.7028 | -100.7474
15.7066 | -101.106
15.7105 | -101.4611
15.7143 | -101.8089
15.7181 | -102.1467
15.7219 | -102.4713
15.7257 | -102.7767
15.7296 | -103.0647
15.7334 | -103.3392
15.7372 | -104.0066
15.741 | -104.8637
15.7448 | -104.9418
15.7487 | -104.9476
15.7525 | -104.9518
15.7563 | -104.9553
15.7601 | -104.9573
15.7639 | -104.9579
15.7678 | -104.9589
15.7716 | -104.9588
15.7754 | -104.957
15.7792 | -104.9542
15.783 | -104.9507
15.7869 | -104.9463
15.7907 | -104.941
15.7945 | -104.9305
15.7983 | -104.9207
15.8021 | -104.9123
15.806 | -104.9034
15.8098 | -104.894
15.8136 | -104.8826
15.8174 | -104.8685
15.8212 | -104.8536
15.8251 | -104.8389
15.8289 | -104.8232
15.8327 | -104.8077
15.8365 | -104.79
15.8403 | -104.7733
15.8442 | -104.7571
15.848 | -104.7393
15.8518 | -104.7213
15.8556 | -104.7031
15.8594 | -104.6875
15.8633 | -104.6712
15.8671 | -104.6568
15.8709 | -104.6413
15.8747 | -104.6305
15.8785 | -104.6213
15.8824 | -104.6113
15.8862 | -104.5997
15.89 | -104.5908
15.8938 | -104.585
15.8976 | -104.579
15.9015 | -104.5751
15.9053 | -104.5739
15.9091 | -104.5742
15.9129 | -104.5732
15.9167 | -104.5736
15.9206 | -104.5772
15.9244 | -104.5799
15.9282 | -104.5868
15.932 | -104.5944
15.9358 | -104.6016
15.9396 | -104.6075
15.9435 | -104.6135
15.9473 | -104.6195
15.9511 | -104.6263
15.9549 | -104.6326
15.9587 | -104.6399
15.9626 | -104.6449
15.9664 | -104.6509
15.9702 | -104.6535
15.974 | -104.6536
15.9778 | -104.6469
15.9817 | -104.6439
15.9855 | -104.6395
15.9893 | -104.636
15.9931 | -104.6347
15.9969 | -104.6336
16.0008 | -104.6302
16.0046 | -104.6232
16.0084 | -104.6145
16.0122 | -104.6063
16.016 | -104.598
16.0199 | -104.5914
16.0237 | -104.5861
16.0275 | -104.5765
16.0313 | -104.5618
16.0351 | -104.5463
16.039 | -104.5341
16.0428 | -104.5217
16.0466 | -104.5122
16.0504 | -104.5023
16.0542 | -104.4896
16.0581 | -104.4758
16.0619 | -104.4621
16.0657 | -104.4513
16.0695 | -104.4386
16.0733 | -104.4291
16.0772 | -104.4222
16.081 | -104.4137
16.0848 | -104.408
16.0886 | -104.4014
16.0924 | -104.396
16.0963 | -104.3916
16.1001 | -104.3894
16.1039 | -104.3853
16.1077 | -104.3828
16.1115 | -104.3844
16.1154 | -104.3917
16.1192 | -104.3986
16.123 | -104.4041
16.1268 | -104.4112
16.1306 | -104.4187
16.1345 | -104.4267
16.1383 | -104.4319
16.1421 | -104.4404
16.1459 | -104.4527
16.1497 | -104.4655
16.1536 | -104.4805
16.1574 | -104.4943
16.1612 | -104.5088
16.165 | -104.5228
16.1688 | -104.5381
16.1727 | -104.5555
16.1765 | -104.5724
16.1803 | -104.5898
16.1841 | -104.6071
16.1879 | -104.6239
16.1917 | -104.6427
16.1956 | -104.6607
16.1994 | -104.6801
16.2032 | -104.6983
16.207 | -104.7155
16.2108 | -104.7316
16.2147 | -104.7472
16.2185 | -104.7631
16.2223 | -104.7786
16.2261 | -104.7923
16.2299 | -104.8066
16.2338 | -104.821
16.2376 | -104.8371
16.2414 | -104.8529
16.2452 | -104.8653
16.249 | -104.8743
16.2529 | -104.8846
16.2567 | -104.8978
16.2605 | -104.9103
16.2643 | -104.9241
16.2681 | -104.9367
16.272 | -104.9448
16.2758 | -104.9537
16.2796 | -104.9603
16.2834 | -104.9675
16.2872 | -104.9747
16.2911 | -104.9794
16.2949 | -104.9847
16.2987 | -104.987
16.3025 | -104.9914
16.3063 | -104.9914
16.3102 | -104.9875
16.314 | -104.9843
16.3178 | -104.9805
16.3216 | -104.977
16.3254 | -104.97
16.3293 | -104.9628
16.3331 | -104.9566
16.3369 | -104.952
16.3407 | -104.9487
16.3445 | -104.9419
16.3484 | -104.9324
16.3522 | -104.9221
16.356 | -104.9115
16.3598 | -104.9002
16.3636 | -104.8895
16.3675 | -104.879
16.3713 | -104.8708
16.3751 | -104.8622
16.3789 | -104.852
16.3827 | -104.8402
16.3866 | -104.8309
16.3904 | -104.8231
16.3942 | -104.8156
16.398 | -104.8095
16.4018 | -104.8046
16.4057 | -104.7951
16.4095 | -104.7878
16.4133 | -104.7833
16.4171 | -104.7777
16.4209 | -104.7735
16.4248 | -104.7699
16.4286 | -104.7671
16.4324 | -104.7627
16.4362 | -104.7604
16.44 | -104.7603
16.4439 | -104.7585
16.4477 | -104.7569
16.4515 | -104.7539
16.4553 | -104.7524
16.4591 | -104.7534
16.4629 | -104.7542
16.4668 | -104.7557
16.4706 | -104.7559
16.4744 | -104.758
16.4782 | -104.7617
16.482 | -104.7679
16.4859 | -104.7703
16.4897 | -104.7701
16.4935 | -104.7709
16.4973 | -104.7717
16.5011 | -104.7742
16.505 | -104.7788
16.5088 | -104.7826
16.5126 | -104.7882
16.5164 | -104.7936
16.5202 | -104.8008
16.5241 | -104.8095
16.5279 | -104.8152
16.5317 | -104.8262
16.5355 | -104.8364
16.5393 | -104.846
16.5432 | -104.8522
16.547 | -104.8589
16.5508 | -104.865
16.5546 | -104.8693
16.5584 | -104.8778
16.5623 | -104.8886
16.5661 | -104.8949
16.5699 | -104.9011
16.5737 | -104.9075
16.5775 | -104.9112
16.5814 | -104.9141
16.5852 | -104.9172
16.589 | -104.9211
16.5928 | -104.9221
16.5966 | -104.9237
16.6005 | -104.9238
16.6043 | -104.9232
16.6081 | -104.9204
16.6119 | -104.9158
16.6157 | -104.9104
16.6196 | -104.9047
16.6234 | -104.8967
16.6272 | -104.8879
16.631 | -104.878
16.6348 | -104.8683
16.6387 | -104.8591
16.6425 | -104.8512
16.6463 | -104.8405
16.6501 | -104.8273
16.6539 | -104.8169
16.6578 | -104.8074
16.6616 | -104.8002
16.6654 | -104.7922
16.6692 | -104.7842
16.673 | -104.7767
16.6769 | -104.7676
16.6807 | -104.761
16.6845 | -104.7551
16.6883 | -104.75
16.6921 | -104.7427
16.696 | -104.7355
16.6998 | -104.7299
16.7036 | -104.7244
16.7074 | -104.7209
16.7112 | -104.7177
16.7151 | -104.7138
16.7189 | -104.7075
16.7227 | -104.704
16.7265 | -104.7053
16.7303 | -104.7066
16.7341 | -104.7073
16.738 | -104.7073
16.7418 | -104.705
16.7456 | -104.7054
16.7494 | -104.7061
16.7532 | -104.7076
16.7571 | -104.7123
16.7609 | -104.7163
16.7647 | -104.7201
16.7685 | -104.725
16.7723 | -104.7321
16.7762 | -104.7386
16.78 | -104.7439
16.7838 | -104.756
16.7876 | -104.7689
16.7914 | -104.7819
16.7953 | -104.795
16.7991 | -104.8053
16.8029 | -104.8155
16.8067 | -104.8321
16.8105 | -104.8472
16.8144 | -104.8605
16.8182 | -104.8751
16.822 | -104.8924
16.8258 | -104.9109
16.8296 | -104.929
16.8335 | -104.9489
16.8373 | -104.9665
16.8411 | -104.983
16.8449 | -104.9978
16.8487 | -105.0139
16.8526 | -105.0305
16.8564 | -105.0448
16.8602 | -105.0604
16.864 | -105.0767
16.8678 | -105.0909
16.8717 | -105.105
16.8755 | -105.1162
16.8793 | -105.1191
16.8831 | -105.1196
16.8869 | -105.1275
16.8908 | -105.1336
16.8946 | -105.135
16.8984 | -105.1357
16.9022 | -105.1326
16.906 | -105.1284
16.9099 | -105.1249
16.9137 | -105.1195
16.9175 | -105.1137
16.9213 | -105.1072
16.9251 | -105.1006
16.929 | -105.0951
16.9328 | -105.0887
16.9366 | -105.0812
16.9404 | -105.0723
16.9442 | -105.0638
16.9481 | -105.0536
16.9519 | -105.0433
16.9557 | -105.0324
16.9595 | -105.0267
16.9633 | -105.0186
16.9672 | -105.0084
16.971 | -105.0004
16.9748 | -104.9937
16.9786 | -104.9863
16.9824 | -104.9833
16.9862 | -104.9792
16.9901 | -104.9762
16.9939 | -104.9744
16.9977 | -104.9727
17.0015 | -104.9745
17.0053 | -104.9804
17.0092 | -104.9879
17.013 | -104.9961
17.0168 | -105.0051
17.0206 | -105.0167
17.0244 | -105.0286
17.0283 | -105.0427
17.0321 | -105.0568
17.0359 | -105.0739
17.0397 | -105.093
17.0435 | -105.1129
17.0474 | -105.1298
17.0512 | -105.1472
17.055 | -105.1672
17.0588 | -105.1885
17.0626 | -105.2092
17.0665 | -105.2274
17.0703 | -105.2442
17.0741 | -105.2603
17.0779 | -105.2776
17.0817 | -105.2885
17.0856 | -105.3038
17.0894 | -105.3198
17.0932 | -105.3328
17.097 | -105.3409
17.1008 | -105.3502
17.1047 | -105.3568
17.1085 | -105.3632
17.1123 | -105.3701
17.1161 | -105.3741
17.1199 | -105.3788
17.1238 | -105.3875
17.1276 | -105.3927
17.1314 | -105.3887
17.1352 | -105.3856
17.139 | -105.3835
17.1429 | -105.3782
17.1467 | -105.3733
17.1505 | -105.3673
17.1543 | -105.3596
17.1581 | -105.3517
17.162 | -105.3442
17.1658 | -105.3342
17.1696 | -105.3236
17.1734 | -105.3099
17.1772 | -105.2976
17.1811 | -105.2861
17.1849 | -105.2755
17.1887 | -105.2664
17.1925 | -105.2584
17.1963 | -105.2501
17.2002 | -105.2443
17.204 | -105.2366
17.2078 | -105.2317
17.2116 | -105.2283
17.2154 | -105.2263
17.2193 | -105.225
17.2231 | -105.2258
17.2269 | -105.2278
17.2307 | -105.2306
17.2345 | -105.2304
17.2384 | -105.2314
17.2422 | -105.2355
17.246 | -105.2368
17.2498 | -105.2391
17.2536 | -105.2418
17.2574 | -105.246
17.2613 | -105.2527
17.2651 | -105.2617
17.2689 | -105.2696
17.2727 | -105.2775
17.2765 | -105.2831
17.2804 | -105.2867
17.2842 | -105.2902
17.288 | -105.2922
17.2918 | -105.2935
17.2956 | -105.2952
17.2995 | -105.295
17.3033 | -105.293
17.3071 | -105.2887
17.3109 | -105.285
17.3147 | -105.2817
17.3186 | -105.2773
17.3224 | -105.2729
17.3262 | -105.2674
17.33 | -105.2618
17.3338 | -105.2558
17.3377 | -105.2487
17.3415 | -105.2427
17.3453 | -105.2385
17.3491 | -105.2338
17.3529 | -105.2295
17.3568 | -105.2238
17.3606 | -105.2193
17.3644 | -105.2177
17.3682 | -105.2173
17.372 | -105.2181
17.3759 | -105.2192
17.3797 | -105.2197
17.3835 | -105.2222
17.3873 | -105.2266
17.3911 | -105.229
17.395 | -105.2322
17.3988 | -105.2378
17.4026 | -105.2433
17.4064 | -105.2471
17.4102 | -105.249
17.4141 | -105.2517
17.4179 | -105.2542
17.4217 | -105.259
17.4255 | -105.2636
17.4293 | -105.2662
17.4332 | -105.2693
17.437 | -105.2725
17.4408 | -105.2739
17.4446 | -105.2737
17.4484 | -105.2732
17.4523 | -105.2731
17.4561 | -105.2727
17.4599 | -105.2724
17.4637 | -105.2706
17.4675 | -105.2661
17.4714 | -105.2641
17.4752 | -105.2618
17.479 | -105.2608
17.4828 | -105.2638
17.4866 | -105.2651
17.4905 | -105.2619
17.4943 | -105.2615
17.4981 | -105.2589
17.5019 | -105.2557
17.5057 | -105.2532
17.5095 | -105.2499
17.5134 | -105.2483
17.5172 | -105.2465
17.521 | -105.2428
17.5248 | -105.2418
17.5286 | -105.2414
17.5325 | -105.2401
17.5363 | -105.2419
17.5401 | -105.2435
17.5439 | -105.2442
17.5477 | -105.2451
17.5516 | -105.2465
17.5554 | -105.2481
17.5592 | -105.249
17.563 | -105.2487
17.5668 | -105.2488
17.5707 | -105.2482
17.5745 | -105.2481
17.5783 | -105.2478
17.5821 | -105.2451
17.5859 | -105.2426
17.5898 | -105.24
17.5936 | -105.238
17.5974 | -105.2367
17.6012 | -105.2364
17.605 | -105.2358
17.6089 | -105.2318
17.6127 | -105.228
17.6165 | -105.2261
17.6203 | -105.2244
17.6241 | -105.2256
17.628 | -105.2276
17.6318 | -105.2325
17.6356 | -105.2349
17.6394 | -105.2377
17.6432 | -105.244
17.6471 | -105.2509
17.6509 | -105.2564
17.6547 | -105.2632
17.6585 | -105.2717
17.6623 | -105.2794
17.6662 | -105.2857
17.67 | -105.2939
17.6738 | -105.3016
17.6776 | -105.3094
17.6814 | -105.3157
17.6853 | -105.321
17.6891 | -105.3265
17.6929 | -105.3344
17.6967 | -105.3429
17.7005 | -105.3477
17.7044 | -105.3516
17.7082 | -105.3559
17.712 | -105.3593
17.7158 | -105.3638
17.7196 | -105.3706
17.7235 | -105.3739
17.7273 | -105.3764
17.7311 | -105.3781
17.7349 | -105.3783
17.7387 | -105.3785
17.7426 | -105.3806
17.7464 | -105.38
17.7502 | -105.3775
17.754 | -105.3749
17.7578 | -105.3704
17.7616 | -105.3646
17.7655 | -105.3572
17.7693 | -105.3474
17.7731 | -105.3412
17.7769 | -105.335
17.7807 | -105.3308
17.7846 | -105.3231
17.7884 | -105.3141
17.7922 | -105.3066
17.796 | -105.2989
17.7998 | -105.2904
17.8037 | -105.2816
17.8075 | -105.2727
17.8113 | -105.2618
17.8151 | -105.2529
17.8189 | -105.2459
17.8228 | -105.2392
17.8266 | -105.2332
17.8304 | -105.2266
17.8342 | -105.2224
17.838 | -105.2191
17.8419 | -105.2158
17.8457 | -105.2128
17.8495 | -105.2083
17.8533 | -105.2056
17.8571 | -105.2036
17.861 | -105.2046
17.8648 | -105.2058
17.8686 | -105.2046
17.8724 | -105.1994
17.8762 | -105.1961
17.8801 | -105.1946
17.8839 | -105.1926
17.8877 | -105.1902
17.8915 | -105.1867
17.8953 | -105.1843
17.8992 | -105.1836
17.903 | -105.1831
17.9068 | -105.1816
17.9106 | -105.1797
17.9144 | -105.1767
17.9183 | -105.1749
17.9221 | -105.174
17.9259 | -105.1706
17.9297 | -105.1671
17.9335 | -105.1641
17.9374 | -105.1621
17.9412 | -105.1593
17.945 | -105.1543
17.9488 | -105.1488
17.9526 | -105.142
17.9565 | -105.1368
17.9603 | -105.1335
17.9641 | -105.126
17.9679 | -105.1157
17.9717 | -105.1056
17.9756 | -105.0967
17.9794 | -105.088
17.9832 | -105.08
17.987 | -105.0723
17.9908 | -105.0636
17.9947 | -105.055
17.9985 | -105.0449
18.0023 | -105.039
18.0061 | -105.0341
18.0099 | -105.0238
18.0138 | -105.0135
18.0176 | -105.0068
18.0214 | -105.0007
18.0252 | -104.9977
18.029 | -104.9926
18.0328 | -104.9873
18.0367 | -104.9794
18.0405 | -104.9728
18.0443 | -104.9687
18.0481 | -104.9668
18.0519 | -104.9648
18.0558 | -104.9607
18.0596 | -104.9557
18.0634 | -104.9523
18.0672 | -104.9498
18.071 | -104.9453
18.0749 | -104.9409
18.0787 | -104.9376
18.0825 | -104.9347
18.0863 | -104.9311
18.0901 | -104.9282
18.094 | -104.9254
18.0978 | -104.9229
18.1016 | -104.9206
18.1054 | -104.9198
18.1092 | -104.9171
18.1131 | -104.9156
18.1169 | -104.9132
18.1207 | -104.9106
18.1245 | -104.9101
18.1283 | -104.9156
18.1322 | -104.9174
18.136 | -104.9174
18.1398 | -104.9186
18.1436 | -104.9193
18.1474 | -104.9186
18.1513 | -104.9178
18.1551 | -104.9174
18.1589 | -104.9148
18.1627 | -104.9125
18.1665 | -104.9083
18.1704 | -104.9048
18.1742 | -104.9018
18.178 | -104.8959
18.1818 | -104.8855
18.1856 | -104.8793
18.1895 | -104.8714
18.1933 | -104.8616
18.1971 | -104.8528
18.2009 | -104.8436
18.2047 | -104.8336
18.2086 | -104.8211
18.2124 | -104.8089
18.2162 | -104.7958
18.22 | -104.7809
18.2238 | -104.7694
18.2277 | -104.7581
18.2315 | -104.7445
18.2353 | -104.7307
18.2391 | -104.7167
18.2429 | -104.7027
18.2468 | -104.6868
18.2506 | -104.6698
18.2544 | -104.6554
18.2582 | -104.6437
18.262 | -104.6274
18.2659 | -104.609
18.2697 | -104.5963
18.2735 | -104.5837
18.2773 | -104.5665
18.2811 | -104.5503
18.2849 | -104.5335
18.2888 | -104.5169
18.2926 | -104.5006
18.2964 | -104.4849
18.3002 | -104.4719
18.304 | -104.4574
18.3079 | -104.4401
18.3117 | -104.4222
18.3155 | -104.4076
18.3193 | -104.3935
18.3231 | -104.3795
18.327 | -104.3646
18.3308 | -104.3513
18.3346 | -104.3357
18.3384 | -104.3196
18.3422 | -104.3047
18.3461 | -104.2932
18.3499 | -104.2798
18.3537 | -104.2659
18.3575 | -104.2527
18.3613 | -104.2401
18.3652 | -104.2278
18.369 | -104.2168
18.3728 | -104.2058
18.3766 | -104.1953
18.3804 | -104.1858
18.3843 | -104.1757
18.3881 | -104.1659
18.3919 | -104.1585
18.3957 | -104.1503
18.3995 | -104.1439
18.4034 | -104.1369
18.4072 | -104.131
18.411 | -104.1249
18.4148 | -104.1197
18.4186 | -104.1186
18.4225 | -104.1185
18.4263 | -104.1205
18.4301 | -104.1234
18.4339 | -104.1267
18.4377 | -104.1334
18.4416 | -104.1402
18.4454 | -104.1481
18.4492 | -104.1543
18.453 | -104.1624
18.4568 | -104.1712
18.4607 | -104.1792
18.4645 | -104.1895
18.4683 | -104.2006
18.4721 | -104.2123
18.4759 | -104.2206
18.4798 | -104.2291
18.4836 | -104.238
18.4874 | -104.246
18.4912 | -104.2535
18.495 | -104.2609
18.4989 | -104.2659
18.5027 | -104.2705
18.5065 | -104.2755
18.5103 | -104.2799
18.5141 | -104.2839
18.518 | -104.2844
18.5218 | -104.2833
18.5256 | -104.2836
18.5294 | -104.2836
18.5332 | -104.2797
18.5371 | -104.2733
18.5409 | -104.2691
18.5447 | -104.262
18.5485 | -104.2527
18.5523 | -104.2425
18.5561 | -104.2331
18.56 | -104.2235
18.5638 | -104.2126
18.5676 | -104.2008
18.5714 | -104.1913
18.5752 | -104.1846
18.5791 | -104.1768
18.5829 | -104.1672
18.5867 | -104.1578
18.5905 | -104.1516
18.5943 | -104.1411
18.5982 | -104.132
18.602 | -104.1247
18.6058 | -104.1156
18.6096 | -104.104
18.6134 | -104.09
18.6173 | -104.0585
18.6211 | -104.0173
18.6249 | -103.9902
18.6287 | -103.9688
18.6325 | -103.9482
18.6364 | -103.9254
18.6402 | -103.9017
18.644 | -103.8739
18.6478 | -103.8456
18.6516 | -103.8136
18.6555 | -103.7819
18.6593 | -103.7514
18.6631 | -103.7158
18.6669 | -103.6808
18.6707 | -103.6447
18.6746 | -103.6078
18.6784 | -103.5682
18.6822 | -103.5244
18.686 | -103.4821
18.6898 | -103.4398
18.6937 | -103.3968
18.6975 | -103.3533
18.7013 | -103.3109
18.7051 | -103.272
18.7089 | -103.2337
18.7128 | -103.1995
18.7166 | -103.1669
18.7204 | -103.1363
18.7242 | -103.1067
18.728 | -103.079
18.7319 | -103.0554
18.7357 | -103.0635
18.7395 | -103.0532
18.7433 | -103.0469
18.7471 | -103.0436
18.751 | -103.043
18.7548 | -103.0449
18.7586 | -103.0534
18.7624 | -103.0647
18.7662 | -103.0813
18.7701 | -103.1011
18.7739 | -103.1233
18.7777 | -103.1501
18.7815 | -103.1795
18.7853 | -103.2122
18.7892 | -103.247
18.793 | -103.2835
18.7968 | -103.3213
18.8006 | -103.36
18.8044 | -103.3987
18.8083 | -103.4366
18.8121 | -103.4733
18.8159 | -103.5118
18.8197 | -103.5515
18.8235 | -103.5902
18.8273 | -103.6272
18.8312 | -103.6602
18.835 | -103.6916
18.8388 | -103.7188
18.8426 | -103.743
18.8464 | -103.765
18.8503 | -103.7859
18.8541 | -103.8028
18.8579 | -103.8176
18.8617 | -103.8319
18.8655 | -103.844
18.8694 | -103.8519
18.8732 | -103.8605
18.877 | -103.8665
18.8808 | -103.8737
18.8846 | -103.8931
18.8885 | -103.9134
18.8923 | -103.915
18.8961 | -103.9097
18.8999 | -103.9022
18.9037 | -103.8955
18.9076 | -103.89
18.9114 | -103.888
18.9152 | -103.8855
18.919 | -103.8844
18.9228 | -103.8821
18.9267 | -103.8793
18.9305 | -103.8805
18.9343 | -103.8808
18.9381 | -103.8816
18.9419 | -103.883
18.9458 | -103.8869
18.9496 | -103.8944
18.9534 | -103.8995
18.9572 | -103.9042
18.961 | -103.9094
18.9649 | -103.9157
18.9687 | -103.9229
18.9725 | -103.9317
18.9763 | -103.9426
18.9801 | -103.9548
18.984 | -103.9679
18.9878 | -103.9811
18.9916 | -103.9968
18.9954 | -104.0131
18.9992 | -104.0288
19.0031 | -104.0412
19.0069 | -104.0541
19.0107 | -104.0671
19.0145 | -104.0828
19.0183 | -104.0988
19.0222 | -104.1144
19.026 | -104.1322
19.0298 | -104.1461
19.0336 | -104.1623
19.0374 | -104.1789
19.0413 | -104.1967
19.0451 | -104.2145
19.0489 | -104.2314
19.0527 | -104.2481
19.0565 | -104.2636
19.0604 | -104.2772
19.0642 | -104.2922
19.068 | -104.3063
19.0718 | -104.3199
19.0756 | -104.3327
19.0795 | -104.3439
19.0833 | -104.3558
19.0871 | -104.3648
19.0909 | -104.3704
19.0947 | -104.3741
19.0985 | -104.377
19.1024 | -104.3818
19.1062 | -104.3864
19.11 | -104.3913
19.1138 | -104.3929
19.1176 | -104.3924
19.1215 | -104.3922
19.1253 | -104.3927
19.1291 | -104.3904
19.1329 | -104.3876
19.1367 | -104.3843
19.1406 | -104.3808
19.1444 | -104.373
19.1482 | -104.3673
19.152 | -104.3601
19.1558 | -104.3511
19.1597 | -104.3421
19.1635 | -104.3306
19.1673 | -104.3183
19.1711 | -104.3051
19.1749 | -104.2907
19.1788 | -104.2738
19.1826 | -104.2561
19.1864 | -104.2384
19.1902 | -104.2184
19.194 | -104.1991
19.1979 | -104.1801
19.2017 | -104.1594
19.2055 | -104.1405
19.2093 | -104.1224
19.2131 | -104.1042
19.217 | -104.0851
19.2208 | -104.0645
19.2246 | -104.0463
19.2284 | -104.0275
19.2322 | -104.0106
19.2361 | -103.9949
19.2399 | -103.9766
19.2437 | -103.9543
19.2475 | -103.9353
19.2513 | -103.9192
19.2552 | -103.9029
19.259 | -103.8905
19.2628 | -103.8769
19.2666 | -103.8645
19.2704 | -103.8529
19.2743 | -103.8427
19.2781 | -103.8312
19.2819 | -103.8223
19.2857 | -103.8225
19.2895 | -103.8161
19.2934 | -103.8118
19.2972 | -103.8088
19.301 | -103.8058
19.3048 | -103.8053
19.3086 | -103.8043
19.3125 | -103.8022
19.3163 | -103.8018
19.3201 | -103.8019
19.3239 | -103.8047
19.3277 | -103.8093
19.3316 | -103.8125
19.3354 | -103.8145
19.3392 | -103.8158
19.343 | -103.8176
19.3468 | -103.8213
19.3506 | -103.8252
19.3545 | -103.8261
19.3583 | -103.8247
19.3621 | -103.8253
19.3659 | -103.8265
19.3697 | -103.8271
19.3736 | -103.8257
19.3774 | -103.823
19.3812 | -103.8187
19.385 | -103.8133
19.3888 | -103.808
19.3927 | -103.8016
19.3965 | -103.7947
19.4003 | -103.7861
19.4041 | -103.7767
19.4079 | -103.7652
19.4118 | -103.7516
19.4156 | -103.7376
19.4194 | -103.7243
19.4232 | -103.71
19.427 | -103.6938
19.4309 | -103.6757
19.4347 | -103.6564
19.4385 | -103.6373
19.4423 | -103.6188
19.4461 | -103.6017
19.45 | -103.5851
19.4538 | -103.5691
19.4576 | -103.5525
19.4614 | -103.5357
19.4652 | -103.5196
19.4691 | -103.5042
19.4729 | -103.4901
19.4767 | -103.4779
19.4805 | -103.4645
19.4843 | -103.4489
19.4882 | -103.4365
19.492 | -103.4254
19.4958 | -103.4181
19.4996 | -103.412
19.5034 | -103.4072
19.5073 | -103.405
19.5111 | -103.402
19.5149 | -103.4006
19.5187 | -103.4018
19.5225 | -103.4021
19.5264 | -103.4043
19.5302 | -103.41
19.534 | -103.4158
19.5378 | -103.4244
19.5416 | -103.4323
19.5455 | -103.4412
19.5493 | -103.4503
19.5531 | -103.4604
19.5569 | -103.4689
19.5607 | -103.4773
19.5646 | -103.4859
19.5684 | -103.4942
19.5722 | -103.5019
19.576 | -103.5113
19.5798 | -103.5215
19.5837 | -103.5307
19.5875 | -103.5393
19.5913 | -103.5487
19.5951 | -103.5562
19.5989 | -103.5637
19.6028 | -103.5711
19.6066 | -103.5789
19.6104 | -103.5859
19.6142 | -103.5935
19.618 | -103.603
19.6218 | -103.6126
19.6257 | -103.6179
19.6295 | -103.6239
19.6333 | -103.6324
19.6371 | -103.6418
19.6409 | -103.6478
19.6448 | -103.6512
19.6486 | -103.654
19.6524 | -103.6608
19.6562 | -103.6679
19.66 | -103.6747
19.6639 | -103.6791
19.6677 | -103.6856
19.6715 | -103.6937
19.6753 | -103.7004
19.6791 | -103.7069
19.683 | -103.7131
19.6868 | -103.7214
19.6906 | -103.7285
19.6944 | -103.7336
19.6982 | -103.7402
19.7021 | -103.7469
19.7059 | -103.7541
19.7097 | -103.7597
19.7135 | -103.7634
19.7173 | -103.7692
19.7212 | -103.7699
19.725 | -103.7662
19.7288 | -103.7667
19.7326 | -103.7672
19.7364 | -103.7656
19.7403 | -103.7642
19.7441 | -103.7647
19.7479 | -103.7625
19.7517 | -103.7584
19.7555 | -103.753
19.7594 | -103.749
19.7632 | -103.747
19.767 | -103.7409
19.7708 | -103.7333
19.7746 | -103.7236
19.7785 | -103.7144
19.7823 | -103.7044
19.7861 | -103.6889
19.7899 | -103.6741
19.7937 | -103.6594
19.7976 | -103.6445
19.8014 | -103.6278
19.8052 | -103.6073
19.809 | -103.5879
19.8128 | -103.5692
19.8167 | -103.5499
19.8205 | -103.5312
19.8243 | -103.5121
19.8281 | -103.4946
19.8319 | -103.4775
19.8358 | -103.4651
19.8396 | -103.4495
19.8434 | -103.434
19.8472 | -103.4207
19.851 | -103.4094
19.8549 | -103.3962
19.8587 | -103.3819
19.8625 | -103.371
19.8663 | -103.3629
19.8701 | -103.3573
19.874 | -103.3525
19.8778 | -103.3497
19.8816 | -103.3499
19.8854 | -103.3518
19.8892 | -103.3514
19.893 | -103.3561
19.8969 | -103.3624
19.9007 | -103.3664
19.9045 | -103.3726
19.9083 | -103.3814
19.9121 | -103.3917
19.916 | -103.4026
19.9198 | -103.4155
19.9236 | -103.4279
19.9274 | -103.4424
19.9312 | -103.4586
19.9351 | -103.4734
19.9389 | -103.4889
19.9427 | -103.5043
19.9465 | -103.5225
19.9503 | -103.5386
19.9542 | -103.5537
19.958 | -103.5681
19.9618 | -103.5835
19.9656 | -103.5993
19.9694 | -103.6133
19.9733 | -103.6284
19.9771 | -103.6461
19.9809 | -103.661
19.9847 | -103.6728
19.9885 | -103.6857
19.9924 | -103.6994
19.9962 | -103.7142
20 | -103.7234
20.0038 | -103.7335
20.0076 | -103.7452
20.0115 | -103.7526
20.0153 | -103.7591
20.0191 | -103.7663
20.0229 | -103.7711
20.0267 | -103.7733
20.0306 | -103.7756
20.0344 | -103.7773
20.0382 | -103.7798
20.042 | -103.7812
20.0458 | -103.7793
20.0497 | -103.7765
20.0535 | -103.7769
20.0573 | -103.7754
20.0611 | -103.7714
20.0649 | -103.7662
20.0688 | -103.7632
20.0726 | -103.7615
20.0764 | -103.7576
20.0802 | -103.7523
20.084 | -103.7484
20.0879 | -103.7426
20.0917 | -103.7346
20.0955 | -103.7229
20.0993 | -103.7123
20.1031 | -103.7028
20.107 | -103.6936
20.1108 | -103.6856
20.1146 | -103.6776
20.1184 | -103.6688
20.1222 | -103.66
20.126 | -103.6526
20.1299 | -103.6444
20.1337 | -103.6354
20.1375 | -103.6275
20.1413 | -103.6226
20.1451 | -103.6193
20.149 | -103.6128
20.1528 | -103.606
20.1566 | -103.5986
20.1604 | -103.5904
20.1642 | -103.5843
20.1681 | -103.579
20.1719 | -103.5745
20.1757 | -103.5734
20.1795 | -103.5701
20.1833 | -103.5672
20.1872 | -103.5663
20.191 | -103.5646
20.1948 | -103.5631
20.1986 | -103.5619
20.2024 | -103.5606
20.2063 | -103.558
20.2101 | -103.5595
20.2139 | -103.5618
20.2177 | -103.5628
20.2215 | -103.5651
20.2254 | -103.5659
20.2292 | -103.5685
20.233 | -103.5728
20.2368 | -103.5768
20.2406 | -103.5819
20.2445 | -103.588
20.2483 | -103.5941
20.2521 | -103.5997
20.2559 | -103.6046
20.2597 | -103.6106
20.2636 | -103.6156
20.2674 | -103.6206
20.2712 | -103.6267
20.275 | -103.6318
20.2788 | -103.6379
20.2827 | -103.6423
20.2865 | -103.6481
20.2903 | -103.6543
20.2941 | -103.659
20.2979 | -103.6648
20.3018 | -103.6686
20.3056 | -103.6719
20.3094 | -103.6761
20.3132 | -103.6806
20.317 | -103.6847
20.3209 | -103.6882
20.3247 | -103.69
20.3285 | -103.6896
20.3323 | -103.6915
20.3361 | -103.6924
20.34 | -103.6915
20.3438 | -103.693
20.3476 | -103.6936
20.3514 | -103.6964
20.3552 | -103.6974
20.3591 | -103.6981
20.3629 | -103.6994
20.3667 | -103.701
20.3705 | -103.7022
20.3743 | -103.7051
20.3782 | -103.7077
20.382 | -103.7131
20.3858 | -103.7175
20.3896 | -103.7222
20.3934 | -103.7239
20.3972 | -103.7259
20.4011 | -103.7261
20.4049 | -103.7269
20.4087 | -103.7273
20.4125 | -103.7267
20.4163 | -103.7262
20.4202 | -103.7292
20.424 | -103.7317
20.4278 | -103.7325
20.4316 | -103.7335
20.4354 | -103.7346
20.4393 | -103.7372
20.4431 | -103.7382
20.4469 | -103.7414
20.4507 | -103.7455
20.4545 | -103.7463
20.4584 | -103.7447
20.4622 | -103.7423
20.466 | -103.7436
20.4698 | -103.7459
20.4736 | -103.7475
20.4775 | -103.7493
20.4813 | -103.7525
20.4851 | -103.7552
20.4889 | -103.7574
20.4927 | -103.7617
20.4966 | -103.7651
20.5004 | -103.7662
20.5042 | -103.7676
20.508 | -103.773
20.5118 | -103.7797
20.5157 | -103.7875
20.5195 | -103.7942
20.5233 | -103.7979
20.5271 | -103.803
20.5309 | -103.8079
20.5348 | -103.8132
20.5386 | -103.8191
20.5424 | -103.8281
20.5462 | -103.8398
20.55 | -103.8494
20.5539 | -103.8571
20.5577 | -103.8622
20.5615 | -103.87
20.5653 | -103.8796
20.5691 | -103.8859
20.573 | -103.8929
20.5768 | -103.8991
20.5806 | -103.9033
20.5844 | -103.9096
20.5882 | -103.9122
20.5921 | -103.913
20.5959 | -103.9122
20.5997 | -103.9135
20.6035 | -103.9144
20.6073 | -103.9124
20.6112 | -103.9088
20.615 | -103.9066
20.6188 | -103.9071
20.6226 | -103.9056
20.6264 | -103.9027
20.6303 | -103.8975
20.6341 | -103.89
20.6379 | -103.8823
20.6417 | -103.8711
20.6455 | -103.8624
20.6494 | -103.8519
20.6532 | -103.843
20.657 | -103.8331
20.6608 | -103.8195
20.6646 | -103.8064
20.6684 | -103.7946
20.6723 | -103.7816
20.6761 | -103.7686
20.6799 | -103.7541
20.6837 | -103.7414
20.6875 | -103.7292
20.6914 | -103.7127
20.6952 | -103.6952
20.699 | -103.6775
20.7028 | -103.6596
20.7066 | -103.6456
20.7105 | -103.6312
20.7143 | -103.6141
20.7181 | -103.5958
20.7219 | -103.578
20.7257 | -103.5609
20.7296 | -103.5418
20.7334 | -103.5227
20.7372 | -103.506
20.741 | -103.4869
20.7448 | -103.4693
20.7487 | -103.4533
20.7525 | -103.4363
20.7563 | -103.4189
20.7601 | -103.4009
20.7639 | -103.384
20.7678 | -103.367
20.7716 | -103.3512
20.7754 | -103.3362
20.7792 | -103.3227
20.783 | -103.3103
20.7869 | -103.2981
20.7907 | -103.2875
20.7945 | -103.2766
20.7983 | -103.2656
20.8021 | -103.2538
20.806 | -103.2412
20.8098 | -103.2306
20.8136 | -103.2235
20.8174 | -103.2167
20.8212 | -103.211
20.8251 | -103.2054
20.8289 | -103.2013
20.8327 | -103.1993
20.8365 | -103.1964
20.8403 | -103.1943
20.8442 | -103.1938
20.848 | -103.1954
20.8518 | -103.1997
20.8556 | -103.201
20.8594 | -103.2035
20.8633 | -103.2082
20.8671 | -103.2136
20.8709 | -103.221
20.8747 | -103.2313
20.8785 | -103.2394
20.8824 | -103.246
20.8862 | -103.2549
20.89 | -103.2649
20.8938 | -103.2724
20.8976 | -103.2803
20.9015 | -103.29
20.9053 | -103.2995
20.9091 | -103.3078
20.9129 | -103.3143
20.9167 | -103.3224
20.9205 | -103.3298
20.9244 | -103.336
20.9282 | -103.3396
20.932 | -103.346
20.9358 | -103.3527
20.9396 | -103.3523
20.9435 | -103.3567
20.9473 | -103.3606
20.9511 | -103.363
20.9549 | -103.3647
20.9587 | -103.3694
20.9626 | -103.3715
20.9664 | -103.3723
20.9702 | -103.3725
20.974 | -103.3719
20.9778 | -103.373
20.9817 | -103.3733
20.9855 | -103.3731
20.9893 | -103.3736
20.9931 | -103.3747
20.9969 | -103.3756
21.0008 | -103.3766
21.0046 | -103.3744
21.0084 | -103.371
21.0122 | -103.3679
21.016 | -103.3662
21.0199 | -103.3659
21.0237 | -103.3648
21.0275 | -103.3641
21.0313 | -103.3606
21.0351 | -103.3595
21.039 | -103.3601
21.0428 | -103.3605
21.0466 | -103.361
21.0504 | -103.3633
21.0542 | -103.3651
21.0581 | -103.3652
21.0619 | -103.3656
21.0657 | -103.3686
21.0695 | -103.3698
21.0733 | -103.3723
21.0772 | -103.375
21.081 | -103.3788
21.0848 | -103.384
21.0886 | -103.3892
21.0924 | -103.3956
21.0963 | -103.4018
21.1001 | -103.4073
21.1039 | -103.4138
21.1077 | -103.4192
21.1115 | -103.4245
21.1154 | -103.4322
21.1192 | -103.4381
21.123 | -103.4432
21.1268 | -103.4461
21.1306 | -103.4477
21.1345 | -103.4502
21.1383 | -103.4524
21.1421 | -103.4548
21.1459 | -103.4568
21.1497 | -103.4594
21.1536 | -103.4617
21.1574 | -103.4618
21.1612 | -103.4621
21.165 | -103.463
21.1688 | -103.462
21.1727 | -103.4592
21.1765 | -103.4574
21.1803 | -103.4565
21.1841 | -103.4551
21.1879 | -103.4513
21.1917 | -103.4484
21.1956 | -103.4445
21.1994 | -103.4411
21.2032 | -103.4367
21.207 | -103.4316
21.2108 | -103.4267
21.2147 | -103.4201
21.2185 | -103.4155
21.2223 | -103.4115
21.2261 | -103.4089
21.2299 | -103.4032
21.2338 | -103.3968
21.2376 | -103.3912
21.2414 | -103.3852
21.2452 | -103.3775
21.249 | -103.3699
21.2529 | -103.3616
21.2567 | -103.3546
21.2605 | -103.3456
21.2643 | -103.338
21.2681 | -103.3322
21.272 | -103.3245
21.2758 | -103.3163
21.2796 | -103.3067
21.2834 | -103.2982
21.2872 | -103.2924
21.2911 | -103.2858
21.2949 | -103.2781
21.2987 | -103.2712
21.3025 | -103.2659
21.3063 | -103.2607
21.3102 | -103.2544
21.314 | -103.2508
21.3178 | -103.248
21.3216 | -103.2435
21.3254 | -103.2399
21.3293 | -103.239
21.3331 | -103.2404
21.3369 | -103.2392
21.3407 | -103.2384
21.3445 | -103.2363
21.3484 | -103.2339
21.3522 | -103.2318
21.356 | -103.2327
21.3598 | -103.2342
21.3636 | -103.236
21.3675 | -103.2376
21.3713 | -103.2406
21.3751 | -103.2434
21.3789 | -103.2465
21.3827 | -103.2477
21.3866 | -103.2495
21.3904 | -103.2567
21.3942 | -103.2657
21.398 | -103.2749
21.4018 | -103.2836
21.4057 | -103.2946
21.4095 | -103.306
21.4133 | -103.315
21.4171 | -103.3245
21.4209 | -103.3355
21.4248 | -103.348
21.4286 | -103.3627
21.4324 | -103.3742
21.4362 | -103.3879
21.44 | -103.4025
21.4439 | -103.4143
21.4477 | -103.4267
21.4515 | -103.4374
21.4553 | -103.4492
21.4591 | -103.462
21.4629 | -103.474
21.4668 | -103.4836
21.4706 | -103.494
21.4744 | -103.5038
21.4782 | -103.5131
21.482 | -103.5233
21.4859 | -103.5303
21.4897 | -103.5299
21.4935 | -103.5373
21.4973 | -103.5427
21.5011 | -103.5476
21.505 | -103.5492
21.5088 | -103.5502
21.5126 | -103.5479
21.5164 | -103.5445
21.5202 | -103.542
21.5241 | -103.5385
21.5279 | -103.5299
21.5317 | -103.5164
21.5355 | -103.505
21.5393 | -103.4981
21.5432 | -103.4901
21.547 | -103.4805
21.5508 | -103.4723
21.5546 | -103.4624
21.5584 | -103.4539
21.5623 | -103.4436
21.5661 | -103.4323
21.5699 | -103.4204
21.5737 | -103.4061
21.5775 | -103.3956
21.5814 | -103.3864
21.5852 | -103.3764
21.589 | -103.3634
21.5928 | -103.3506
21.5966 | -103.3397
21.6005 | -103.3271
21.6043 | -103.3124
21.6081 | -103.2977
21.6119 | -103.2831
21.6157 | -103.2725
21.6196 | -103.2586
21.6234 | -103.2428
21.6272 | -103.2283
21.631 | -103.2147
21.6348 | -103.1993
21.6387 | -103.1856
21.6425 | -103.1733
21.6463 | -103.1628
21.6501 | -103.1519
21.6539 | -103.1424
21.6578 | -103.135
21.6616 | -103.1246
21.6654 | -103.114
21.6692 | -103.1049
21.673 | -103.0978
21.6769 | -103.0903
21.6807 | -103.0829
21.6845 | -103.0761
21.6883 | -103.071
21.6921 | -103.0652
21.696 | -103.055
21.6998 | -103.047
21.7036 | -103.04
21.7074 | -103.035
21.7112 | -103.0276
21.7151 | -103.0194
21.7189 | -103.0124
21.7227 | -103.0028
21.7265 | -102.9949
21.7303 | -102.9894
21.7341 | -102.9817
21.738 | -102.9744
21.7418 | -102.9644
21.7456 | -102.9546
21.7494 | -102.9447
21.7532 | -102.9321
21.7571 | -102.921
21.7609 | -102.9087
21.7647 | -102.8942
21.7685 | -102.8789
21.7723 | -102.8656
21.7762 | -102.8518
21.78 | -102.8368
21.7838 | -102.819
21.7876 | -102.7994
21.7914 | -102.7807
21.7953 | -102.7641
21.7991 | -102.747
21.8029 | -102.7291
21.8067 | -102.7122
21.8105 | -102.6946
21.8144 | -102.6775
21.8182 | -102.6607
21.822 | -102.6428
21.8258 | -102.6216
21.8296 | -102.604
21.8335 | -102.5862
21.8373 | -102.5701
21.8411 | -102.5545
21.8449 | -102.5329
21.8487 | -102.5086
21.8526 | -102.4859
21.8564 | -102.4634
21.8602 | -102.4407
21.864 | -102.4199
21.8678 | -102.3976
21.8717 | -102.3757
21.8755 | -102.3545
21.8793 | -102.332
21.8831 | -102.3107
21.8869 | -102.2882
21.8908 | -102.2628
21.8946 | -102.2379
21.8984 | -102.2142
21.9022 | -102.1907
21.906 | -102.1689
21.9099 | -102.1498
21.9137 | -102.1313
21.9175 | -102.1108
21.9213 | -102.0899
21.9251 | -102.068
21.929 | -102.0495
21.9328 | -102.0319
21.9366 | -102.0144
21.9404 | -101.9981
21.9442 | -101.9816
21.9481 | -101.9642
21.9519 | -101.9491
21.9557 | -101.9331
21.9595 | -101.9191
21.9633 | -101.9066
21.9672 | -101.8921
21.971 | -101.8794
21.9748 | -101.8677
21.9786 | -101.8572
21.9824 | -101.8476
21.9862 | -101.8417
21.9901 | -101.8369
21.9939 | -101.8338
21.9977 | -101.8322
22.0015 | -101.8301
22.0053 | -101.8283
22.0092 | -101.8246
22.013 | -101.8234
22.0168 | -101.8243
22.0206 | -101.8268
22.0244 | -101.8299
22.0283 | -101.8335
22.0321 | -101.836
22.0359 | -101.8395
22.0397 | -101.8489
22.0435 | -101.8525
22.0474 | -101.8567
22.0512 | -101.8635
22.055 | -101.8699
22.0588 | -101.8761
22.0626 | -101.8843
22.0665 | -101.8931
22.0703 | -101.8992
22.0741 | -101.9065
22.0779 | -101.9129
22.0817 | -101.9166
22.0856 | -101.9191
22.0894 | -101.9221
22.0932 | -101.9262
22.097 | -101.9316
22.1008 | -101.9356
22.1047 | -101.9372
22.1085 | -101.9369
22.1123 | -101.9363
22.1161 | -101.9357
22.1199 | -101.9317
22.1238 | -101.9258
22.1276 | -101.9177
22.1314 | -101.9084
22.1352 | -101.901
22.139 | -101.8954
22.1429 | -101.8837
22.1467 | -101.8711
22.1505 | -101.8568
22.1543 | -101.8452
22.1581 | -101.8374
22.162 | -101.8258
22.1658 | -101.8125
22.1696 | -101.7992
22.1734 | -101.7826
22.1772 | -101.7666
22.1811 | -101.7508
22.1849 | -101.7339
22.1887 | -101.7166
22.1925 | -101.6984
22.1963 | -101.6798
22.2002 | -101.6602
22.204 | -101.6383
22.2078 | -101.6171
22.2116 | -101.597
22.2154 | -101.5778
22.2193 | -101.5582
22.2231 | -101.5387
22.2269 | -101.5191
22.2307 | -101.5008
22.2345 | -101.482
22.2384 | -101.4635
22.2422 | -101.4432
22.246 | -101.4236
22.2498 | -101.4049
22.2536 | -101.3837
22.2574 | -101.3638
22.2613 | -101.3409
22.2651 | -101.3172
22.2689 | -101.2972
22.2727 | -101.2767
22.2765 | -101.2546
22.2804 | -101.2336
22.2842 | -101.2109
22.288 | -101.189
22.2918 | -101.166
22.2956 | -101.1439
22.2995 | -101.1213
22.3033 | -101.0985
22.3071 | -101.0756
22.3109 | -101.0527
22.3147 | -101.0329
22.3186 | -101.0136
22.3224 | -100.9934
22.3262 | -100.9731
22.33 | -100.9513
22.3338 | -100.9296
22.3377 | -100.9096
22.3415 | -100.8873
22.3453 | -100.8681
22.3491 | -100.8466
22.3529 | -100.8213
22.3568 | -100.7948
22.3606 | -100.7718
22.3644 | -100.7499
22.3682 | -100.7258
22.372 | -100.7036
22.3759 | -100.6813
22.3797 | -100.6571
22.3835 | -100.6345
22.3873 | -100.6155
22.3911 | -100.5946
22.395 | -100.5719
22.3988 | -100.5535
22.4026 | -100.5352
22.4064 | -100.5173
22.4102 | -100.4986
22.4141 | -100.4799
22.4179 | -100.4625
22.4217 | -100.4458
22.4255 | -100.4285
22.4293 | -100.4087
22.4332 | -100.3908
22.437 | -100.3761
22.4408 | -100.3621
22.4446 | -100.3497
22.4484 | -100.3363
22.4523 | -100.3226
22.4561 | -100.3074
22.4599 | -100.2918
22.4637 | -100.2767
22.4675 | -100.2631
22.4714 | -100.2484
22.4752 | -100.2312
22.479 | -100.2152
22.4828 | -100.2003
22.4866 | -100.1866
22.4905 | -100.1751
22.4943 | -100.163
22.4981 | -100.1533
22.5019 | -100.1452
22.5057 | -100.1364
22.5095 | -100.129
22.5134 | -100.1216
22.5172 | -100.1157
22.521 | -100.1123
22.5248 | -100.1095
22.5286 | -100.1063
22.5325 | -100.1021
22.5363 | -100.0988
22.5401 | -100.0968
22.5439 | -100.0959
22.5477 | -100.095
22.5516 | -100.0975
22.5554 | -100.1018
22.5592 | -100.1058
22.563 | -100.1114
22.5668 | -100.1177
22.5707 | -100.1257
22.5745 | -100.1341
22.5783 | -100.1446
22.5821 | -100.154
22.5859 | -100.162
22.5898 | -100.1718
22.5936 | -100.1865
22.5974 | -100.1971
22.6012 | -100.2087
22.605 | -100.2217
22.6089 | -100.2314
22.6127 | -100.2413
22.6165 | -100.2498
22.6203 | -100.2584
22.6241 | -100.2689
22.628 | -100.2786
22.6318 | -100.2869
22.6356 | -100.2946
22.6394 | -100.3034
22.6432 | -100.3133
22.6471 | -100.3213
22.6509 | -100.3315
22.6547 | -100.3427
22.6585 | -100.353
22.6623 | -100.365
22.6662 | -100.3756
22.67 | -100.3854
22.6738 | -100.3984
22.6776 | -100.4118
22.6814 | -100.4222
22.6853 | -100.43
22.6891 | -100.4391
22.6929 | -100.4487
22.6967 | -100.4597
22.7005 | -100.4711
22.7044 | -100.4784
22.7082 | -100.4829
22.712 | -100.488
22.7158 | -100.4976
22.7196 | -100.5037
22.7235 | -100.5064
22.7273 | -100.5094
22.7311 | -100.5131
22.7349 | -100.5169
22.7387 | -100.5218
22.7426 | -100.525
22.7464 | -100.5286
22.7502 | -100.5333
22.754 | -100.5366
22.7578 | -100.5387
22.7616 | -100.5403
22.7655 | -100.5405
22.7693 | -100.541
22.7731 | -100.5415
22.7769 | -100.5431
22.7807 | -100.5454
22.7846 | -100.547
22.7884 | -100.5467
22.7922 | -100.5459
22.796 | -100.5474
22.7998 | -100.5486
22.8037 | -100.548
22.8075 | -100.5464
22.8113 | -100.5444
22.8151 | -100.5427
22.8189 | -100.5411
22.8228 | -100.5404
22.8266 | -100.541
22.8304 | -100.5391
22.8342 | -100.5379
22.838 | -100.5377
22.8419 | -100.5387
22.8457 | -100.5407
22.8495 | -100.5424
22.8533 | -100.5432
22.8571 | -100.5446
22.861 | -100.5484
22.8648 | -100.5519
22.8686 | -100.5575
22.8724 | -100.5628
22.8762 | -100.5669
22.8801 | -100.5739
22.8839 | -100.5819
22.8877 | -100.5895
22.8915 | -100.5971
22.8953 | -100.6065
22.8992 | -100.6176
22.903 | -100.6292
22.9068 | -100.6411
22.9106 | -100.6515
22.9144 | -100.6613
22.9183 | -100.671
22.9221 | -100.6823
22.9259 | -100.6936
22.9297 | -100.7026
22.9335 | -100.7151
22.9374 | -100.7261
22.9412 | -100.7384
22.945 | -100.7507
22.9488 | -100.7618
22.9526 | -100.7738
22.9565 | -100.7857
22.9603 | -100.7958
22.9641 | -100.8031
22.9679 | -100.8093
22.9717 | -100.8172
22.9756 | -100.822
22.9794 | -100.8261
22.9832 | -100.8289
22.987 | -100.8312
22.9908 | -100.8342
22.9947 | -100.8381
22.9985 | -100.8397
23.0023 | -100.8404
23.0061 | -100.8421
23.0099 | -100.8445
23.0138 | -100.8456
23.0176 | -100.8411
23.0214 | -100.8361
23.0252 | -100.8334
23.029 | -100.8314
23.0328 | -100.8277
23.0367 | -100.8229
23.0405 | -100.814
23.0443 | -100.8039
23.0481 | -100.797
23.0519 | -100.7895
23.0558 | -100.7798
23.0596 | -100.7677
23.0634 | -100.7514
23.0672 | -100.7374
23.071 | -100.7242
23.0749 | -100.7091
23.0787 | -100.6918
23.0825 | -100.6753
23.0863 | -100.6595
23.0901 | -100.6441
23.094 | -100.6273
23.0978 | -100.609
23.1016 | -100.5909
23.1054 | -100.5739
23.1092 | -100.5568
23.1131 | -100.5396
23.1169 | -100.5221
23.1207 | -100.5047
23.1245 | -100.4874
23.1283 | -100.472
23.1322 | -100.4565
23.136 | -100.4405
23.1398 | -100.4246
23.1436 | -100.4132
23.1474 | -100.3997
23.1513 | -100.3861
23.1551 | -100.3744
23.1589 | -100.3665
23.1627 | -100.3583
23.1665 | -100.3505
23.1704 | -100.3412
23.1742 | -100.331
23.178 | -100.324
23.1818 | -100.3185
23.1856 | -100.3117
23.1895 | -100.3056
23.1933 | -100.3023
23.1971 | -100.2992
23.2009 | -100.2957
23.2047 | -100.2937
23.2086 | -100.2935
23.2124 | -100.2931
23.2162 | -100.2954
23.22 | -100.2982
23.2238 | -100.2998
23.2277 | -100.3025
23.2315 | -100.307
23.2353 | -100.3128
23.2391 | -100.3186
23.2429 | -100.3254
23.2468 | -100.335
23.2506 | -100.3454
23.2544 | -100.3563
23.2582 | -100.369
23.262 | -100.3812
23.2659 | -100.3928
23.2697 | -100.4028
23.2735 | -100.4134
23.2773 | -100.4239
23.2811 | -100.4356
23.2849 | -100.4483
23.2888 | -100.462
23.2926 | -100.4777
23.2964 | -100.492
23.3002 | -100.5051
23.304 | -100.5198
23.3079 | -100.533
23.3117 | -100.547
23.3155 | -100.562
23.3193 | -100.5767
23.3231 | -100.5904
23.327 | -100.6033
23.3308 | -100.6148
23.3346 | -100.6258
23.3384 | -100.6388
23.3422 | -100.6527
23.3461 | -100.6677
23.3499 | -100.6818
23.3537 | -100.6948
23.3575 | -100.709
23.3613 | -100.7242
23.3652 | -100.7393
23.369 | -100.7542
23.3728 | -100.7681
23.3766 | -100.7811
23.3804 | -100.7928
23.3843 | -100.8051
23.3881 | -100.8169
23.3919 | -100.8327
23.3957 | -100.8484
23.3995 | -100.8615
23.4034 | -100.8746
23.4072 | -100.888
23.411 | -100.9006
23.4148 | -100.9132
23.4186 | -100.9259
23.4225 | -100.9374
23.4263 | -100.9469
23.4301 | -100.9561
23.4339 | -100.9632
23.4377 | -100.9719
23.4416 | -100.9824
23.4454 | -100.9911
23.4492 | -100.9974
23.453 | -101.0047
23.4568 | -101.0108
23.4607 | -101.0173
23.4645 | -101.0244
23.4683 | -101.0306
23.4721 | -101.035
23.4759 | -101.0384
23.4798 | -101.0412
23.4836 | -101.0433
23.4874 | -101.0432
23.4912 | -101.044
23.495 | -101.0452
23.4989 | -101.0462
23.5027 | -101.049
23.5065 | -101.052
23.5103 | -101.0541
23.5141 | -101.0572
23.518 | -101.0612
23.5218 | -101.0621
23.5256 | -101.0626
23.5294 | -101.0643
23.5332 | -101.0672
23.5371 | -101.07
23.5409 | -101.0697
23.5447 | -101.0715
23.5485 | -101.073
23.5523 | -101.0764
23.5561 | -101.0824
23.56 | -101.0864
23.5638 | -101.0911
23.5676 | -101.097
23.5714 | -101.1048
23.5752 | -101.1091
23.5791 | -101.1119
23.5829 | -101.1143
23.5867 | -101.1174
23.5905 | -101.1217
23.5943 | -101.127
23.5982 | -101.1344
23.602 | -101.1419
23.6058 | -101.1485
23.6096 | -101.1526
23.6134 | -101.1565
23.6173 | -101.1619
23.6211 | -101.1661
23.6249 | -101.1695
23.6287 | -101.1725
23.6325 | -101.1749
23.6364 | -101.1773
23.6402 | -101.1801
23.644 | -101.1825
23.6478 | -101.1822
23.6516 | -101.1828
23.6555 | -101.1839
23.6593 | -101.1866
23.6631 | -101.1899
23.6669 | -101.1898
23.6707 | -101.188
23.6746 | -101.1868
23.6784 | -101.1861
23.6822 | -101.183
23.686 | -101.1802
23.6898 | -101.1789
23.6937 | -101.1756
23.6975 | -101.1745
23.7013 | -101.1738
23.7051 | -101.1723
23.7089 | -101.1719
23.7128 | -101.1702
23.7166 | -101.1677
23.7204 | -101.1661
23.7242 | -101.1654
23.728 | -101.1633
23.7319 | -101.1597
23.7357 | -101.1582
23.7395 | -101.1589
23.7433 | -101.1604
23.7471 | -101.1613
23.751 | -101.1605
23.7548 | -101.1614
23.7586 | -101.1658
23.7624 | -101.1716
23.7662 | -101.1757
23.7701 | -101.1793
23.7739 | -101.1818
23.7777 | -101.1849
23.7815 | -101.1899
23.7853 | -101.1958
23.7892 | -101.2001
23.793 | -101.2045
23.7968 | -101.2105
23.8006 | -101.2161
23.8044 | -101.2225
23.8083 | -101.2313
23.8121 | -101.2404
23.8159 | -101.2482
23.8197 | -101.257
23.8235 | -101.2657
23.8273 | -101.276
23.8312 | -101.2865
23.835 | -101.2968
23.8388 | -101.3085
23.8426 | -101.3196
23.8464 | -101.3295
23.8503 | -101.3404
23.8541 | -101.3547
23.8579 | -101.3686
23.8617 | -101.3819
23.8655 | -101.3955
23.8694 | -101.4081
23.8732 | -101.4215
23.877 | -101.4371
23.8808 | -101.45
23.8846 | -101.4602
23.8885 | -101.4718
23.8923 | -101.4822
23.8961 | -101.494
23.8999 | -101.5036
23.9037 | -101.5112
23.9076 | -101.5195
23.9114 | -101.5285
23.9152 | -101.5368
23.919 | -101.5458
23.9228 | -101.5522
23.9267 | -101.5565
23.9305 | -101.5595
23.9343 | -101.5596
23.9381 | -101.56
23.9419 | -101.5595
23.9458 | -101.5619
23.9496 | -101.5625
23.9534 | -101.5598
23.9572 | -101.5582
23.961 | -101.5564
23.9649 | -101.5553
23.9687 | -101.5528
23.9725 | -101.5478
23.9763 | -101.5418
23.9801 | -101.5353
23.984 | -101.5257
23.9878 | -101.5135
23.9916 | -101.5035
23.9954 | -101.4938
23.9992 | -101.4816
24.0031 | -101.4707
24.0069 | -101.4606
24.0107 | -101.4486
24.0145 | -101.4369
24.0183 | -101.427
24.0222 | -101.4162
24.026 | -101.4028
24.0298 | -101.3919
24.0336 | -101.3802
24.0374 | -101.3658
24.0413 | -101.3536
24.0451 | -101.3436
24.0489 | -101.3339
24.0527 | -101.3257
24.0565 | -101.3153
24.0604 | -101.3049
24.0642 | -101.2975
24.068 | -101.2876
24.0718 | -101.2783
24.0756 | -101.2684
24.0795 | -101.2589
24.0833 | -101.2504
24.0871 | -101.2429
24.0909 | -101.2376
24.0947 | -101.2326
24.0985 | -101.2272
24.1024 | -101.2222
24.1062 | -101.2181
24.11 | -101.2172
24.1138 | -101.2163
24.1176 | -101.2149
24.1215 | -101.2134
24.1253 | -101.2111
24.1291 | -101.2123
24.1329 | -101.2141
24.1367 | -101.2157
24.1406 | -101.217
24.1444 | -101.2165
24.1482 | -101.2151
24.152 | -101.2169
24.1558 | -101.2194
24.1597 | -101.221
24.1635 | -101.2246
24.1673 | -101.2249
24.1711 | -101.2263
24.1749 | -101.2262
24.1788 | -101.224
24.1826 | -101.2239
24.1864 | -101.2243
24.1902 | -101.2222
24.194 | -101.2182
24.1979 | -101.2126
24.2017 | -101.2055
24.2055 | -101.2
24.2093 | -101.1947
24.2131 | -101.1892
24.217 | -101.1804
24.2208 | -101.1708
24.2246 | -101.1632
24.2284 | -101.153
24.2322 | -101.1405
24.2361 | -101.1252
24.2399 | -101.1117
24.2437 | -101.0936
24.2475 | -101.0815
24.2513 | -101.0681
24.2552 | -101.0538
24.259 | -101.0399
24.2628 | -101.0226
24.2666 | -101.0055
24.2704 | -100.9889
24.2743 | -100.9712
24.2781 | -100.9552
24.2819 | -100.9381
24.2857 | -100.9213
24.2895 | -100.9022
24.2934 | -100.8832
24.2972 | -100.8663
24.301 | -100.8474
24.3048 | -100.8276
24.3086 | -100.8094
24.3125 | -100.7931
24.3163 | -100.7769
24.3201 | -100.7616
24.3239 | -100.7461
24.3277 | -100.7316
24.3316 | -100.7184
24.3354 | -100.7047
24.3392 | -100.6907
24.343 | -100.6799
24.3468 | -100.6699
24.3506 | -100.6584
24.3545 | -100.6467
24.3583 | -100.6352
24.3621 | -100.6227
24.3659 | -100.6122
24.3697 | -100.6027
24.3736 | -100.5926
24.3774 | -100.5847
24.3812 | -100.5779
24.385 | -100.5691
24.3888 | -100.5611
24.3927 | -100.5519
24.3965 | -100.5408
24.4003 | -100.534
24.4041 | -100.5275
24.4079 | -100.5204
24.4118 | -100.5137
24.4156 | -100.5047
24.4194 | -100.4962
24.4232 | -100.4897
24.427 | -100.4827
24.4309 | -100.474
24.4347 | -100.467
24.4385 | -100.4588
24.4423 | -100.4507
24.4461 | -100.4443
24.45 | -100.4375
24.4538 | -100.429
24.4576 | -100.4196
24.4614 | -100.4099
24.4652 | -100.402
24.4691 | -100.396
24.4729 | -100.3882
24.4767 | -100.3813
24.4805 | -100.3755
24.4843 | -100.3695
24.4882 | -100.364
24.492 | -100.3592
24.4958 | -100.3516
24.4996 | -100.3453
24.5034 | -100.3415
24.5073 | -100.3379
24.5111 | -100.3343
24.5149 | -100.3293
24.5187 | -100.3256
24.5225 | -100.3236
24.5264 | -100.3212
24.5302 | -100.3201
24.534 | -100.3206
24.5378 | -100.3206
24.5416 | -100.3242
24.5455 | -100.3281
24.5493 | -100.3302
24.5531 | -100.3328
24.5569 | -100.3341
24.5607 | -100.3358
24.5646 | -100.3396
24.5684 | -100.344
24.5722 | -100.3483
24.576 | -100.3521
24.5798 | -100.3551
24.5837 | -100.3498
24.5875 | -100.3341
24.5913 | -100.3103
24.5951 | -100.2938
24.5989 | -100.279
24.6028 | -100.2656
24.6066 | -100.2488
24.6104 | -100.2271
24.6142 | -100.2053
24.618 | -100.1861
24.6218 | -100.1609
24.6257 | -100.1335
24.6295 | -100.1079
24.6333 | -100.079
24.6371 | -100.049
24.6409 | -100.0169
24.6448 | -99.9817
24.6486 | -99.9449
24.6524 | -99.9058
24.6562 | -99.8636
24.66 | -99.8188
24.6639 | -99.7723
24.6677 | -99.7243
24.6715 | -99.676
24.6753 | -99.6282
24.6791 | -99.5768
24.683 | -99.5228
24.6868 | -99.4679
24.6906 | -99.413
24.6944 | -99.357
24.6982 | -99.3019
24.7021 | -99.2419
24.7059 | -99.1825
24.7097 | -99.1229
24.7135 | -99.0654
24.7173 | -99.0133
24.7212 | -98.9636
24.725 | -98.9149
24.7288 | -98.8674
24.7326 | -98.8227
24.7364 | -98.7802
24.7403 | -98.7423
24.7441 | -98.7082
24.7479 | -98.6767
24.7517 | -98.6471
24.7555 | -98.6203
24.7594 | -98.5981
24.7632 | -98.5801
24.767 | -98.5664
24.7708 | -98.5555
24.7746 | -98.5481
24.7785 | -98.5443
24.7823 | -98.5438
24.7861 | -98.5466
24.7899 | -98.5524
24.7937 | -98.5611
24.7976 | -98.5989
24.8014 | -98.6136
24.8052 | -98.6323
24.809 | -98.654
24.8128 | -98.6785
24.8167 | -98.7044
24.8205 | -98.7312
24.8243 | -98.76
24.8281 | -98.7916
24.8319 | -98.8249
24.8358 | -98.8618
24.8396 | -98.9015
24.8434 | -98.9411
24.8472 | -98.9821
24.851 | -99.0248
24.8549 | -99.0677
24.8587 | -99.1107
24.8625 | -99.1543
24.8663 | -99.2
24.8701 | -99.2466
24.874 | -99.2927
24.8778 | -99.3381
24.8816 | -99.382
24.8854 | -99.4277
24.8892 | -99.4741
24.893 | -99.5192
24.8969 | -99.5651
24.9007 | -99.6114
24.9045 | -99.6574
24.9083 | -99.7033
24.9121 | -99.7489
24.916 | -99.7936
24.9198 | -99.8379
24.9236 | -99.8828
24.9274 | -99.9297
24.9312 | -99.977
24.9351 | -100.0287
24.9389 | -100.0852
24.9427 | -100.1433
24.9465 | -100.1906
24.9503 | -100.2341
24.9542 | -100.2749
24.958 | -100.3164
24.9618 | -100.3594
24.9656 | -100.3996
24.9694 | -100.4336
24.9733 | -100.4687
24.9771 | -100.5027
24.9809 | -100.5333
24.9847 | -100.5639
24.9885 | -100.5936
24.9924 | -100.6229
24.9962 | -100.652
25 | -100.6815
25.0038 | -100.7113
25.0076 | -100.7399
25.0115 | -100.767
25.0153 | -100.795
25.0191 | -100.8223
25.0229 | -100.8486
25.0267 | -100.8729
25.0306 | -100.8977
25.0344 | -100.9228
25.0382 | -100.9481
25.042 | -100.9715
25.0458 | -100.9968
25.0497 | -101.0209
25.0535 | -101.0514
25.0573 | -101.0797
25.0611 | -101.1082
25.0649 | -101.1324
25.0688 | -101.1493
25.0726 | -101.1648
25.0764 | -101.1805
25.0802 | -101.1978
25.084 | -101.2144
25.0879 | -101.2291
25.0917 | -101.242
25.0955 | -101.254
25.0993 | -101.264
25.1031 | -101.2769
25.107 | -101.2902
25.1108 | -101.301
25.1146 | -101.312
25.1184 | -101.3202
25.1222 | -101.3283
25.126 | -101.337
25.1299 | -101.3464
25.1337 | -101.3554
25.1375 | -101.3644
25.1413 | -101.3739
25.1451 | -101.3825
25.149 | -101.39
25.1528 | -101.3962
25.1566 | -101.4015
25.1604 | -101.4045
25.1642 | -101.4081
25.1681 | -101.412
25.1719 | -101.4148
25.1757 | -101.4185
25.1795 | -101.424
25.1833 | -101.4266
25.1872 | -101.429
25.191 | -101.4292
25.1948 | -101.4291
25.1986 | -101.4301
25.2024 | -101.4308
25.2063 | -101.4301
25.2101 | -101.429
25.2139 | -101.4304
25.2177 | -101.4322
25.2215 | -101.4316
25.2254 | -101.4292
25.2292 | -101.4273
25.233 | -101.4269
25.2368 | -101.4292
25.2406 | -101.4277
25.2445 | -101.4249
25.2483 | -101.4217
25.2521 | -101.4199
25.2559 | -101.417
25.2597 | -101.414
25.2636 | -101.4117
25.2674 | -101.4098
25.2712 | -101.4069
25.275 | -101.4019
25.2788 | -101.3964
25.2827 | -101.39
25.2865 | -101.3853
25.2903 | -101.3808
25.2941 | -101.3775
25.2979 | -101.3768
25.3018 | -101.3756
25.3056 | -101.3731
25.3094 | -101.3709
25.3132 | -101.3668
25.317 | -101.364
25.3209 | -101.3617
25.3247 | -101.3593
25.3285 | -101.3583
25.3323 | -101.3575
25.3361 | -101.3549
25.34 | -101.352
25.3438 | -101.3497
25.3476 | -101.3487
25.3514 | -101.3477
25.3552 | -101.3474
25.3591 | -101.3462
25.3629 | -101.346
25.3667 | -101.3458
25.3705 | -101.3457
25.3743 | -101.3463
25.3782 | -101.3477
25.382 | -101.3486
25.3858 | -101.349
25.3896 | -101.3504
25.3934 | -101.351
25.3972 | -101.3526
25.4011 | -101.3537
25.4049 | -101.3519
25.4087 | -101.3491
25.4125 | -101.349
25.4163 | -101.3495
25.4202 | -101.3491
25.424 | -101.3468
25.4278 | -101.3432
25.4316 | -101.3408
25.4354 | -101.3384
25.4393 | -101.3358
25.4431 | -101.3335
25.4469 | -101.3296
25.4507 | -101.3258
25.4545 | -101.3217
25.4584 | -101.3189
25.4622 | -101.3132
25.466 | -101.3071
25.4698 | -101.3021
25.4736 | -101.2978
25.4775 | -101.2933
25.4813 | -101.2873
25.4851 | -101.2799
25.4889 | -101.2716
25.4927 | -101.2626
25.4966 | -101.2537
25.5004 | -101.2449
25.5042 | -101.2374
25.508 | -101.2298
25.5118 | -101.2237
25.5157 | -101.2175
25.5195 | -101.2118
25.5233 | -101.2059
25.5271 | -101.1989
25.5309 | -101.1925
25.5348 | -101.1866
25.5386 | -101.1794
25.5424 | -101.1726
25.5462 | -101.1658
25.55 | -101.1602
25.5539 | -101.1529
25.5577 | -101.1473
25.5615 | -101.1433
25.5653 | -101.1403
25.5691 | -101.1366
25.573 | -101.1292
25.5768 | -101.1236
25.5806 | -101.1203
25.5844 | -101.1196
25.5882 | -101.118
25.5921 | -101.1163
25.5959 | -101.1155
25.5997 | -101.1167
25.6035 | -101.1185
25.6073 | -101.1206
25.6112 | -101.1234
25.615 | -101.1271
25.6188 | -101.1304
25.6226 | -101.136
25.6264 | -101.1431
25.6303 | -101.1488
25.6341 | -101.1549
25.6379 | -101.1614
25.6417 | -101.169
25.6455 | -101.1762
25.6494 | -101.1831
25.6532 | -101.1907
25.657 | -101.1979
25.6608 | -101.2056
25.6646 | -101.2153
25.6684 | -101.2251
25.6723 | -101.2353
25.6761 | -101.2461
25.6799 | -101.2561
25.6837 | -101.2661
25.6875 | -101.2767
25.6914 | -101.2854
25.6952 | -101.2936
25.699 | -101.3033
25.7028 | -101.3125
25.7066 | -101.3201
25.7105 | -101.326
25.7143 | -101.3319
25.7181 | -101.338
25.7219 | -101.3449
25.7257 | -101.3511
25.7296 | -101.359
25.7334 | -101.3667
25.7372 | -101.3725
25.741 | -101.3788
25.7448 | -101.3862
25.7487 | -101.3952
25.7525 | -101.4042
25.7563 | -101.4136
25.7601 | -101.4203
25.7639 | -101.4282
25.7678 | -101.439
25.7716 | -101.4503
25.7754 | -101.4601
25.7792 | -101.4682
25.783 | -101.4766
25.7869 | -101.4872
25.7907 | -101.4992
25.7945 | -101.5082
25.7983 | -101.5176
25.8021 | -101.5281
25.806 | -101.5372
25.8098 | -101.5488
25.8136 | -101.56
25.8174 | -101.5692
25.8212 | -101.577
25.8251 | -101.5878
25.8289 | -101.5982
25.8327 | -101.6083
25.8365 | -101.6183
25.8403 | -101.6273
25.8442 | -101.6367
25.848 | -101.6478
25.8518 | -101.6609
25.8556 | -101.6737
25.8594 | -101.6862
25.8633 | -101.6985
25.8671 | -101.7074
25.8709 | -101.7157
25.8747 | -101.724
25.8785 | -101.7318
25.8824 | -101.7391
25.8862 | -101.7476
25.89 | -101.7561
25.8938 | -101.7645
25.8976 | -101.7726
25.9015 | -101.7812
25.9053 | -101.7895
25.9091 | -101.7976
25.9129 | -101.8058
25.9167 | -101.8136
25.9205 | -101.82
25.9244 | -101.8264
25.9282 | -101.8331
25.932 | -101.8399
25.9358 | -101.8468
25.9396 | -101.8557
25.9435 | -101.8638
25.9473 | -101.8686
25.9511 | -101.8699
25.9549 | -101.8737
25.9587 | -101.877
25.9626 | -101.8771
25.9664 | -101.8755
25.9702 | -101.8745
25.974 | -101.8742
25.9778 | -101.8704
25.9817 | -101.8676
25.9855 | -101.8653
25.9893 | -101.8634
25.9931 | -101.861
25.9969 | -101.8595
26.0008 | -101.8569
26.0046 | -101.8537
26.0084 | -101.8461
26.0122 | -101.8369
26.016 | -101.8273
26.0199 | -101.8199
26.0237 | -101.8131
26.0275 | -101.8037
26.0313 | -101.7928
26.0351 | -101.7828
26.039 | -101.7734
26.0428 | -101.7629
26.0466 | -101.7525
26.0504 | -101.7437
26.0542 | -101.7342
26.0581 | -101.7247
26.0619 | -101.7151
26.0657 | -101.7028
26.0695 | -101.6917
26.0733 | -101.6815
26.0772 | -101.672
26.081 | -101.6615
26.0848 | -101.6485
26.0886 | -101.6363
26.0924 | -101.6256
26.0963 | -101.6146
26.1001 | -101.6029
26.1039 | -101.5917
26.1077 | -101.5797
26.1115 | -101.5688
26.1154 | -101.5576
26.1192 | -101.5456
26.123 | -101.536
26.1268 | -101.5262
26.1306 | -101.5172
26.1345 | -101.5073
26.1383 | -101.4975
26.1421 | -101.4889
26.1459 | -101.4819
26.1497 | -101.4751
26.1536 | -101.4688
26.1574 | -101.4628
26.1612 | -101.4582
26.165 | -101.454
26.1688 | -101.4494
26.1727 | -101.4443
26.1765 | -101.4398
26.1803 | -101.4362
26.1841 | -101.4347
26.1879 | -101.4331
26.1917 | -101.4289
26.1956 | -101.425
26.1994 | -101.4242
26.2032 | -101.4212
26.207 | -101.4199
26.2108 | -101.4192
26.2147 | -101.4174
26.2185 | -101.4154
26.2223 | -101.4137
26.2261 | -101.4135
26.2299 | -101.4125
26.2338 | -101.4117
26.2376 | -101.4095
26.2414 | -101.4084
26.2452 | -101.4074
26.249 | -101.4063
26.2529 | -101.4024
26.2567 | -101.3988
26.2605 | -101.3949
26.2643 | -101.3904
26.2681 | -101.3863
26.272 | -101.3794
26.2758 | -101.3713
26.2796 | -101.3632
26.2834 | -101.3564
26.2872 | -101.3498
26.2911 | -101.3447
26.2949 | -101.3373
26.2987 | -101.3275
26.3025 | -101.3181
26.3063 | -101.3098
26.3102 | -101.3017
26.314 | -101.2915
26.3178 | -101.2818
26.3216 | -101.2738
26.3254 | -101.2637
26.3293 | -101.2525
26.3331 | -101.2406
26.3369 | -101.23
26.3407 | -101.2198
26.3445 | -101.2083
26.3484 | -101.1948
26.3522 | -101.1826
26.356 | -101.17
26.3598 | -101.1568
26.3636 | -101.1439
26.3675 | -101.1305
26.3713 | -101.1181
26.3751 | -101.1073
26.3789 | -101.0996
26.3827 | -101.0923
26.3866 | -101.0852
26.3904 | -101.076
26.3942 | -101.0669
26.398 | -101.0603
26.4018 | -101.0545
26.4057 | -101.049
26.4095 | -101.0444
26.4133 | -101.0398
26.4171 | -101.0326
26.4209 | -101.0258
26.4248 | -101.0215
26.4286 | -101.0189
26.4324 | -101.0157
26.4362 | -101.0147
26.44 | -101.0113
26.4439 | -101.0087
26.4477 | -101.0109
26.4515 | -101.0094
26.4553 | -101.0075
26.4591 | -101.008
26.4629 | -101.0083
26.4668 | -101.0073
26.4706 | -101.0061
26.4744 | -101.007
26.4782 | -101.0075
26.482 | -101.0076
26.4859 | -101.0073
26.4897 | -101.0073
26.4935 | -101.0069
26.4973 | -101.0052
26.5011 | -101.0048
26.505 | -101.0037
26.5088 | -101.0014
26.5126 | -100.9996
26.5164 | -100.9957
26.5202 | -100.9919
26.5241 | -100.9883
26.5279 | -100.983
26.5317 | -100.9781
26.5355 | -100.9724
26.5393 | -100.9666
26.5432 | -100.9593
26.547 | -100.9524
26.5508 | -100.9454
26.5546 | -100.9368
26.5584 | -100.9293
26.5623 | -100.9211
26.5661 | -100.913
26.5699 | -100.9045
26.5737 | -100.8908
26.5775 | -100.8792
26.5814 | -100.8686
26.5852 | -100.8561
26.589 | -100.8436
26.5928 | -100.83
26.5966 | -100.8134
26.6005 | -100.7955
26.6043 | -100.777
26.6081 | -100.7599
26.6119 | -100.7426
26.6157 | -100.7226
26.6196 | -100.7038
26.6234 | -100.6859
26.6272 | -100.669
26.631 | -100.6515
26.6348 | -100.633
26.6387 | -100.6126
26.6425 | -100.5913
26.6463 | -100.5715
26.6501 | -100.5532
26.6539 | -100.5356
26.6578 | -100.5132
26.6616 | -100.4907
26.6654 | -100.4694
26.6692 | -100.4486
26.673 | -100.4283
26.6769 | -100.4045
26.6807 | -100.3818
26.6845 | -100.3605
26.6883 | -100.3398
26.6921 | -100.3199
26.696 | -100.2997
26.6998 | -100.2798
26.7036 | -100.2609
26.7074 | -100.2426
26.7112 | -100.2244
26.7151 | -100.205
26.7189 | -100.1877
26.7227 | -100.173
26.7265 | -100.1559
26.7303 | -100.1384
26.7341 | -100.1199
26.738 | -100.1023
26.7418 | -100.0859
26.7456 | -100.0699
26.7494 | -100.0548
26.7532 | -100.0406
26.7571 | -100.0256
26.7609 | -100.0108
26.7647 | -99.9981
26.7685 | -99.9841
26.7723 | -99.972
26.7762 | -99.9622
26.78 | -99.9532
26.7838 | -99.9448
26.7876 | -99.9381
26.7914 | -99.9332
26.7953 | -99.9297
26.7991 | -99.9244
26.8029 | -99.9177
26.8067 | -99.9119
26.8105 | -99.9047
26.8144 | -99.8987
26.8182 | -99.8947
26.822 | -99.8895
26.8258 | -99.881
26.8296 | -99.8735
26.8335 | -99.8684
26.8373 | -99.8635
26.8411 | -99.8591
26.8449 | -99.8558
26.8487 | -99.8523
26.8526 | -99.847
26.8564 | -99.8424
26.8602 | -99.8379
26.864 | -99.8315
26.8678 | -99.8244
26.8717 | -99.8187
26.8755 | -99.8123
26.8793 | -99.8032
26.8831 | -99.7928
26.8869 | -99.7837
26.8908 | -99.7745
26.8946 | -99.7654
26.8984 | -99.7528
26.9022 | -99.739
26.906 | -99.7257
26.9099 | -99.7123
26.9137 | -99.6981
26.9175 | -99.6814
26.9213 | -99.6624
26.9251 | -99.6432
26.929 | -99.6241
26.9328 | -99.6053
26.9366 | -99.5845
26.9404 | -99.5642
26.9442 | -99.5435
26.9481 | -99.5231
26.9519 | -99.5025
26.9557 | -99.4796
26.9595 | -99.4532
26.9633 | -99.4278
26.9672 | -99.4041
26.971 | -99.3807
26.9748 | -99.3588
26.9786 | -99.3354
26.9824 | -99.3138
26.9862 | -99.2924
26.9901 | -99.2711
26.9939 | -99.2491
26.9977 | -99.2275
27.0015 | -99.211
27.0053 | -99.1895
27.0092 | -99.169
27.013 | -99.149
27.0168 | -99.1288
27.0206 | -99.1116
27.0244 | -99.0956
27.0283 | -99.0795
27.0321 | -99.0623
27.0359 | -99.0454
27.0397 | -99.0326
27.0435 | -99.0213
27.0474 | -99.0098
27.0512 | -98.9989
27.055 | -98.9899
27.0588 | -98.9817
27.0626 | -98.9716
27.0665 | -98.964
27.0703 | -98.9578
27.0741 | -98.9519
27.0779 | -98.9464
27.0817 | -98.9398
27.0856 | -98.9343
27.0894 | -98.9312
27.0932 | -98.9278
27.097 | -98.9212
27.1008 | -98.9108
27.1047 | -98.9019
27.1085 | -98.8946
27.1123 | -98.8866
27.1161 | -98.8814
27.1199 | -98.871
27.1238 | -98.8584
27.1276 | -98.8483
27.1314 | -98.8364
27.1352 | -98.8233
27.139 | -98.8121
27.1429 | -98.8001
27.1467 | -98.7865
27.1505 | -98.7739
27.1543 | -98.7591
27.1581 | -98.7422
27.162 | -98.7242
27.1658 | -98.7053
27.1696 | -98.6854
27.1734 | -98.6645
27.1772 | -98.6407
27.1811 | -98.6189
27.1849 | -98.5984
27.1887 | -98.5772
27.1925 | -98.5553
27.1963 | -98.5329
27.2002 | -98.5085
27.204 | -98.4819
27.2078 | -98.4546
27.2116 | -98.4263
27.2154 | -98.3955
27.2193 | -98.3651
27.2231 | -98.3355
27.2269 | -98.3043
27.2307 | -98.2744
27.2345 | -98.2476
27.2384 | -98.22
27.2422 | -98.1924
27.246 | -98.1625
27.2498 | -98.1326
27.2536 | -98.1044
27.2574 | -98.0779
27.2613 | -98.053
27.2651 | -98.0283
27.2689 | -98.0025
27.2727 | -97.9779
27.2765 | -97.9522
27.2804 | -97.9307
27.2842 | -97.9105
27.288 | -97.89
27.2918 | -97.8701
27.2956 | -97.8536
27.2995 | -97.8392
27.3033 | -97.8272
27.3071 | -97.8156
27.3109 | -97.8042
27.3147 | -97.7937
27.3186 | -97.7835
27.3224 | -97.7754
27.3262 | -97.7687
27.33 | -97.7624
27.3338 | -97.7577
27.3377 | -97.754
27.3415 | -97.753
27.3453 | -97.753
27.3491 | -97.7527
27.3529 | -97.7517
27.3568 | -97.7507
27.3606 | -97.7511
27.3644 | -97.7535
27.3682 | -97.7564
27.372 | -97.7609
27.3759 | -97.7668
27.3797 | -97.7704
27.3835 | -97.7754
27.3873 | -97.7816
27.3911 | -97.787
27.395 | -97.7926
27.3988 | -97.7985
27.4026 | -97.8029
27.4064 | -97.8045
27.4102 | -97.8074
27.4141 | -97.8106
27.4179 | -97.8139
27.4217 | -97.8191
27.4255 | -97.8246
27.4293 | -97.8289
27.4332 | -97.8322
27.437 | -97.8365
27.4408 | -97.8402
27.4446 | -97.8452
27.4484 | -97.8511
27.4523 | -97.8543
27.4561 | -97.8578
27.4599 | -97.8618
27.4637 | -97.8656
27.4675 | -97.8693
27.4714 | -97.8698
27.4752 | -97.871
27.479 | -97.8733
27.4828 | -97.8776
27.4866 | -97.8837
27.4905 | -97.8877
27.4943 | -97.8914
27.4981 | -97.8956
27.5019 | -97.8967
27.5057 | -97.8979
27.5095 | -97.8993
27.5134 | -97.8996
27.5172 | -97.9017
27.521 | -97.9027
27.5248 | -97.9009
27.5286 | -97.9007
27.5325 | -97.8992
27.5363 | -97.8987
27.5401 | -97.8973
27.5439 | -97.8948
27.5477 | -97.8927
27.5516 | -97.8941
27.5554 | -97.8937
27.5592 | -97.8933
27.563 | -97.8945
27.5668 | -97.8965
27.5707 | -97.8969
27.5745 | -97.8968
27.5783 | -97.8986
27.5821 | -97.9004
27.5859 | -97.9023
27.5898 | -97.9048
27.5936 | -97.9083
27.5974 | -97.9131
27.6012 | -97.9186
27.605 | -97.9239
27.6089 | -97.9246
27.6127 | -97.9243
27.6165 | -97.9251
27.6203 | -97.9257
27.6241 | -97.9262
27.628 | -97.9272
27.6318 | -97.9291
27.6356 | -97.9347
27.6394 | -97.9413
27.6432 | -97.9449
27.6471 | -97.9475
27.6509 | -97.9479
27.6547 | -97.9496
27.6585 | -97.9537
27.6623 | -97.9577
27.6662 | -97.963
27.67 | -97.9657
27.6738 | -97.9674
27.6776 | -97.9695
27.6814 | -97.9733
27.6853 | -97.9752
27.6891 | -97.9771
27.6929 | -97.9778
27.6967 | -97.9784
27.7005 | -97.9779
27.7044 | -97.9769
27.7082 | -97.9737
27.712 | -97.9688
27.7158 | -97.9663
27.7196 | -97.9657
27.7235 | -97.9655
27.7273 | -97.9648
27.7311 | -97.9634
27.7349 | -97.9618
27.7387 | -97.96
27.7426 | -97.9596
27.7464 | -97.9593
27.7502 | -97.9573
27.754 | -97.9532
27.7578 | -97.9492
27.7616 | -97.9442
27.7655 | -97.9393
27.7693 | -97.9363
27.7731 | -97.9316
27.7769 | -97.9275
27.7807 | -97.9237
27.7846 | -97.9189
27.7884 | -97.9146
27.7922 | -97.9124
27.796 | -97.9085
27.7998 | -97.9006
27.8037 | -97.8943
27.8075 | -97.8878
27.8113 | -97.8802
27.8151 | -97.8722
27.8189 | -97.8645
27.8228 | -97.859
27.8266 | -97.8539
27.8304 | -97.8481
27.8342 | -97.8418
27.838 | -97.8348
27.8419 | -97.8272
27.8457 | -97.8218
27.8495 | -97.8156
27.8533 | -97.8071
27.8571 | -97.7995
27.861 | -97.7947
27.8648 | -97.7894
27.8686 | -97.7838
27.8724 | -97.7783
27.8762 | -97.7723
27.8801 | -97.7662
27.8839 | -97.7598
27.8877 | -97.7531
27.8915 | -97.7468
27.8953 | -97.7419
27.8992 | -97.7379
27.903 | -97.732
27.9068 | -97.7248
27.9106 | -97.7201
27.9144 | -97.7166
27.9183 | -97.7117
27.9221 | -97.7082
27.9259 | -97.7022
27.9297 | -97.6982
27.9335 | -97.6949
27.9374 | -97.6897
27.9412 | -97.6816
27.945 | -97.676
27.9488 | -97.6693
27.9526 | -97.6643
27.9565 | -97.6591
27.9603 | -97.6528
27.9641 | -97.6458
27.9679 | -97.6374
27.9717 | -97.6291
27.9756 | -97.6243
27.9794 | -97.6208
27.9832 | -97.6172
27.987 | -97.6146
27.9908 | -97.6133
27.9947 | -97.6112
27.9985 | -97.6086
28.0023 | -97.6059