## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 218 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 217
Hood | 02-92B/P | 155
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1001 | hPa |  | PRC PRO 01
Humidity | 39.5 | %/Hr
Internal temperature of calibration bench | 23.59 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1003 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.6 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.92 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.02 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3118 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 02/05/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 14/05/2018
Engineer | Anthony HUE | Approval |  | 18/05/2018
CEO | Robert GLEMOT | Approval |  | 18/05/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.356 | 57.6059
0.012209 | -35.9494 | 38.4883
0.018313 | -34.9418 | 28.3765
0.024417 | -34.5004 | 21.8683
0.030521 | -34.3603 | 17.9761
0.036626 | -34.2223 | 14.8883
0.04273 | -34.1458 | 12.6662
0.048834 | -34.0939 | 11.0161
0.054938 | -34.0429 | 9.7479
0.061043 | -33.9886 | 8.6496
0.067147 | -33.9777 | 7.7735
0.073251 | -33.973 | 7.0126
0.079355 | -33.9777 | 6.4561
0.08546 | -33.9861 | 6.0235
0.091564 | -33.9852 | 5.6432
0.097668 | -33.9624 | 5.3147
0.10377 | -33.9279 | 4.924
0.10988 | -33.9067 | 4.5406
0.11598 | -33.9025 | 4.1833
0.12209 | -33.9181 | 3.9136
0.12819 | -33.9453 | 3.6983
0.13429 | -33.9707 | 3.488
0.1404 | -33.9849 | 3.2633
0.1465 | -33.9986 | 3.0454
0.15261 | -34.0069 | 2.8362
0.15871 | -34.0075 | 2.6461
0.16482 | -33.9993 | 2.4651
0.17092 | -33.9791 | 2.3497
0.17702 | -33.9637 | 2.2183
0.18313 | -33.949 | 2.1105
0.18923 | -33.9381 | 2.001
0.19534 | -33.9327 | 1.9091
0.20144 | -33.9326 | 1.8186
0.20754 | -33.9375 | 1.7603
0.21365 | -33.9411 | 1.6967
0.21975 | -33.9467 | 1.6362
0.22586 | -33.9555 | 1.5715
0.23196 | -33.9642 | 1.4935
0.23807 | -33.9722 | 1.4065
0.24417 | -33.9737 | 1.3044
0.25027 | -33.9763 | 1.2126
0.25638 | -33.9762 | 1.1146
0.26248 | -33.9744 | 1.0261
0.26859 | -33.9726 | 0.94205
0.27469 | -33.9727 | 0.85985
0.2808 | -33.9786 | 0.80581
0.2869 | -33.9835 | 0.74546
0.293 | -33.9892 | 0.69531
0.29911 | -33.9958 | 0.66147
0.30521 | -34.002 | 0.62959
0.31132 | -34.0067 | 0.61101
0.31742 | -34.0114 | 0.58566
0.32353 | -34.0158 | 0.5606
0.32963 | -34.0194 | 0.53528
0.33573 | -34.0221 | 0.50704
0.34184 | -34.0252 | 0.47623
0.34794 | -34.0285 | 0.42889
0.35405 | -34.0318 | 0.38658
0.36015 | -34.0351 | 0.34196
0.36626 | -34.0373 | 0.29578
0.37236 | -34.0395 | 0.24308
0.37846 | -34.0411 | 0.18795
0.38457 | -34.0398 | 0.14224
0.39067 | -34.0404 | 0.0982
0.39678 | -34.0387 | 0.05221
0.40288 | -34.0364 | 0.011481
0.40899 | -34.0329 | -0.028385
0.41509 | -34.0292 | -0.068118
0.42119 | -34.0246 | -0.10718
0.4273 | -34.0212 | -0.14646
0.4334 | -34.0183 | -0.1865
0.43951 | -34.0156 | -0.2275
0.44561 | -34.014 | -0.26755
0.45172 | -34.0154 | -0.30446
0.45782 | -34.0153 | -0.34187
0.46392 | -34.0164 | -0.3764
0.47003 | -34.0181 | -0.40947
0.47613 | -34.0204 | -0.43774
0.48224 | -34.023 | -0.46337
0.48834 | -34.0268 | -0.47809
0.49445 | -34.0297 | -0.49643
0.50055 | -34.0334 | -0.51416
0.50665 | -34.0364 | -0.52589
0.51276 | -34.0389 | -0.53905
0.51886 | -34.0414 | -0.54936
0.52497 | -34.042 | -0.55477
0.53107 | -34.0436 | -0.56101
0.53717 | -34.0452 | -0.56874
0.54328 | -34.0463 | -0.57396
0.54938 | -34.0471 | -0.58248
0.55549 | -34.0474 | -0.5928
0.56159 | -34.0473 | -0.61201
0.5677 | -34.0474 | -0.62819
0.5738 | -34.0469 | -0.64651
0.5799 | -34.0471 | -0.66406
0.58601 | -34.0476 | -0.68944
0.59211 | -34.0482 | -0.72334
0.59822 | -34.0486 | -0.74916
0.60432 | -34.0493 | -0.77822
0.61043 | -34.0501 | -0.80941
0.61653 | -34.0512 | -0.83904
0.62263 | -34.0519 | -0.87077
0.62874 | -34.0534 | -0.89765
0.63484 | -34.0547 | -0.92677
0.64095 | -34.0563 | -0.95472
0.64705 | -34.0579 | -0.9819
0.65316 | -34.0594 | -1.0063
0.65926 | -34.061 | -1.0314
0.66536 | -34.0619 | -1.05
0.67147 | -34.0632 | -1.0703
0.67757 | -34.0646 | -1.0913
0.68368 | -34.0656 | -1.1107
0.68978 | -34.067 | -1.1294
0.69589 | -34.068 | -1.1467
0.70199 | -34.0685 | -1.1644
0.70809 | -34.0692 | -1.1833
0.7142 | -34.0699 | -1.2002
0.7203 | -34.0705 | -1.2184
0.72641 | -34.0707 | -1.2352
0.73251 | -34.0716 | -1.2537
0.73862 | -34.072 | -1.2742
0.74472 | -34.0731 | -1.291
0.75082 | -34.0735 | -1.308
0.75693 | -34.0739 | -1.3278
0.76303 | -34.074 | -1.3463
0.76914 | -34.0737 | -1.3704
0.77524 | -34.0737 | -1.3931
0.78135 | -34.0737 | -1.4139
0.78745 | -34.0739 | -1.4365
0.79355 | -34.074 | -1.4588
0.79966 | -34.0741 | -1.4826
0.80576 | -34.0741 | -1.5107
0.81187 | -34.0736 | -1.5367
0.81797 | -34.0733 | -1.564
0.82408 | -34.0732 | -1.5883
0.83018 | -34.0729 | -1.6164
0.83628 | -34.0729 | -1.6457
0.84239 | -34.0724 | -1.6708
0.84849 | -34.0718 | -1.6952
0.8546 | -34.0716 | -1.7205
0.8607 | -34.0713 | -1.7464
0.86681 | -34.071 | -1.7729
0.87291 | -34.0704 | -1.7966
0.87901 | -34.07 | -1.8176
0.88512 | -34.0698 | -1.8393
0.89122 | -34.0697 | -1.861
0.89733 | -34.0697 | -1.8807
0.90343 | -34.0696 | -1.9008
0.90953 | -34.0702 | -1.9164
0.91564 | -34.0701 | -1.9324
0.92174 | -34.0701 | -1.9474
0.92785 | -34.0701 | -1.9619
0.93395 | -34.0706 | -1.9751
0.94006 | -34.071 | -1.9907
0.94616 | -34.0716 | -1.9978
0.95226 | -34.0722 | -2.0113
0.95837 | -34.0728 | -2.024
0.96447 | -34.0736 | -2.0359
0.97058 | -34.0743 | -2.0472
0.97668 | -34.0748 | -2.0564
0.98279 | -34.0756 | -2.0615
0.98889 | -34.0766 | -2.0696
0.99499 | -34.0772 | -2.0793
1.0011 | -34.0782 | -2.0874
1.0072 | -34.0789 | -2.0935
1.0133 | -34.0796 | -2.1005
1.0194 | -34.0805 | -2.1073
1.0255 | -34.081 | -2.115
1.0316 | -34.0814 | -2.1214
1.0377 | -34.0817 | -2.1282
1.0438 | -34.082 | -2.1343
1.0499 | -34.0816 | -2.1403
1.056 | -34.0816 | -2.1454
1.0621 | -34.0815 | -2.1514
1.0682 | -34.0819 | -2.1581
1.0743 | -34.0818 | -2.1653
1.0805 | -34.0821 | -2.1723
1.0866 | -34.0817 | -2.183
1.0927 | -34.0815 | -2.1907
1.0988 | -34.0817 | -2.2011
1.1049 | -34.0814 | -2.2092
1.111 | -34.0809 | -2.2198
1.1171 | -34.0806 | -2.2292
1.1232 | -34.08 | -2.2428
1.1293 | -34.0797 | -2.255
1.1354 | -34.0794 | -2.2668
1.1415 | -34.0789 | -2.28
1.1476 | -34.0786 | -2.2934
1.1537 | -34.078 | -2.3065
1.1598 | -34.0782 | -2.324
1.1659 | -34.0783 | -2.3368
1.172 | -34.0782 | -2.3504
1.1781 | -34.0781 | -2.3647
1.1842 | -34.078 | -2.3805
1.1903 | -34.0788 | -2.3977
1.1964 | -34.0792 | -2.4127
1.2025 | -34.0794 | -2.4278
1.2086 | -34.08 | -2.4444
1.2147 | -34.0806 | -2.4605
1.2209 | -34.0814 | -2.4758
1.227 | -34.083 | -2.4926
1.2331 | -34.084 | -2.5078
1.2392 | -34.0851 | -2.525
1.2453 | -34.0863 | -2.5387
1.2514 | -34.0876 | -2.5536
1.2575 | -34.0891 | -2.5693
1.2636 | -34.0911 | -2.5838
1.2697 | -34.0926 | -2.5984
1.2758 | -34.094 | -2.6141
1.2819 | -34.0955 | -2.6301
1.288 | -34.097 | -2.6457
1.2941 | -34.0987 | -2.6597
1.3002 | -34.1004 | -2.6718
1.3063 | -34.1019 | -2.6849
1.3124 | -34.1035 | -2.6978
1.3185 | -34.1052 | -2.7116
1.3246 | -34.1068 | -2.7253
1.3307 | -34.1085 | -2.7379
1.3368 | -34.1099 | -2.7481
1.3429 | -34.1113 | -2.7599
1.349 | -34.1127 | -2.7712
1.3551 | -34.1141 | -2.7848
1.3613 | -34.1155 | -2.7964
1.3674 | -34.1166 | -2.8065
1.3735 | -34.1178 | -2.8171
1.3796 | -34.119 | -2.8275
1.3857 | -34.12 | -2.8375
1.3918 | -34.1209 | -2.8494
1.3979 | -34.122 | -2.8604
1.404 | -34.1226 | -2.8711
1.4101 | -34.1235 | -2.8821
1.4162 | -34.1244 | -2.8926
1.4223 | -34.1251 | -2.9036
1.4284 | -34.1258 | -2.9138
1.4345 | -34.1265 | -2.9252
1.4406 | -34.1272 | -2.9358
1.4467 | -34.1278 | -2.9458
1.4528 | -34.1285 | -2.9558
1.4589 | -34.1293 | -2.9667
1.465 | -34.1296 | -2.9757
1.4711 | -34.1305 | -2.9853
1.4772 | -34.131 | -2.9963
1.4833 | -34.1317 | -3.008
1.4894 | -34.1324 | -3.0193
1.4955 | -34.1331 | -3.0308
1.5016 | -34.1337 | -3.0424
1.5078 | -34.1343 | -3.0556
1.5139 | -34.1351 | -3.0677
1.52 | -34.1357 | -3.0806
1.5261 | -34.1362 | -3.0919
1.5322 | -34.1366 | -3.1032
1.5383 | -34.1369 | -3.115
1.5444 | -34.1376 | -3.1274
1.5505 | -34.1381 | -3.1395
1.5566 | -34.1386 | -3.1513
1.5627 | -34.1392 | -3.1633
1.5688 | -34.14 | -3.1749
1.5749 | -34.1403 | -3.1868
1.581 | -34.1409 | -3.2027
1.5871 | -34.1415 | -3.2171
1.5932 | -34.142 | -3.2295
1.5993 | -34.1425 | -3.2446
1.6054 | -34.143 | -3.2588
1.6115 | -34.1435 | -3.274
1.6176 | -34.144 | -3.2902
1.6237 | -34.1445 | -3.3038
1.6298 | -34.1449 | -3.3188
1.6359 | -34.1454 | -3.3329
1.642 | -34.1459 | -3.3473
1.6482 | -34.1465 | -3.3642
1.6543 | -34.1472 | -3.3799
1.6604 | -34.1477 | -3.3946
1.6665 | -34.1481 | -3.4098
1.6726 | -34.1485 | -3.4258
1.6787 | -34.1492 | -3.4425
1.6848 | -34.1494 | -3.4598
1.6909 | -34.1499 | -3.476
1.697 | -34.1503 | -3.4923
1.7031 | -34.1507 | -3.5082
1.7092 | -34.151 | -3.5237
1.7153 | -34.1514 | -3.5405
1.7214 | -34.1514 | -3.5571
1.7275 | -34.1516 | -3.5736
1.7336 | -34.1518 | -3.5902
1.7397 | -34.152 | -3.6059
1.7458 | -34.1524 | -3.6231
1.7519 | -34.1525 | -3.64
1.758 | -34.1527 | -3.6555
1.7641 | -34.153 | -3.6708
1.7702 | -34.1532 | -3.6869
1.7763 | -34.1535 | -3.7018
1.7824 | -34.1536 | -3.7185
1.7885 | -34.1538 | -3.7331
1.7947 | -34.1541 | -3.7484
1.8008 | -34.1543 | -3.7623
1.8069 | -34.1545 | -3.777
1.813 | -34.1547 | -3.7915
1.8191 | -34.1548 | -3.8061
1.8252 | -34.1548 | -3.8201
1.8313 | -34.1551 | -3.8349
1.8374 | -34.1553 | -3.849
1.8435 | -34.1556 | -3.8617
1.8496 | -34.1556 | -3.8762
1.8557 | -34.1558 | -3.8907
1.8618 | -34.1559 | -3.9024
1.8679 | -34.156 | -3.9158
1.874 | -34.1562 | -3.9293
1.8801 | -34.1563 | -3.9423
1.8862 | -34.1566 | -3.955
1.8923 | -34.1567 | -3.9674
1.8984 | -34.1566 | -3.9783
1.9045 | -34.1567 | -3.9898
1.9106 | -34.1567 | -4.0021
1.9167 | -34.1566 | -4.0137
1.9228 | -34.1567 | -4.0247
1.9289 | -34.1565 | -4.0339
1.9351 | -34.1567 | -4.0437
1.9412 | -34.1568 | -4.0545
1.9473 | -34.157 | -4.0664
1.9534 | -34.157 | -4.078
1.9595 | -34.157 | -4.0898
1.9656 | -34.1566 | -4.1004
1.9717 | -34.1565 | -4.1112
1.9778 | -34.1565 | -4.1229
1.9839 | -34.1566 | -4.1344
1.99 | -34.1567 | -4.1449
1.9961 | -34.1566 | -4.1555
2.0022 | -34.1566 | -4.1658
2.0083 | -34.1564 | -4.1765
2.0144 | -34.1565 | -4.1866
2.0205 | -34.1566 | -4.1985
2.0266 | -34.1567 | -4.2102
2.0327 | -34.1569 | -4.2217
2.0388 | -34.1571 | -4.2355
2.0449 | -34.1573 | -4.2462
2.051 | -34.1574 | -4.2563
2.0571 | -34.1573 | -4.2679
2.0632 | -34.1573 | -4.2789
2.0693 | -34.1574 | -4.2902
2.0754 | -34.1577 | -4.3022
2.0816 | -34.1579 | -4.3141
2.0877 | -34.1579 | -4.3256
2.0938 | -34.1582 | -4.3377
2.0999 | -34.1583 | -4.3497
2.106 | -34.1587 | -4.3619
2.1121 | -34.1588 | -4.3743
2.1182 | -34.159 | -4.3849
2.1243 | -34.1592 | -4.396
2.1304 | -34.1594 | -4.4067
2.1365 | -34.1598 | -4.4183
2.1426 | -34.16 | -4.4318
2.1487 | -34.1605 | -4.4438
2.1548 | -34.161 | -4.4567
2.1609 | -34.1616 | -4.4687
2.167 | -34.162 | -4.4815
2.1731 | -34.1624 | -4.4932
2.1792 | -34.1632 | -4.5049
2.1853 | -34.1637 | -4.5176
2.1914 | -34.1643 | -4.53
2.1975 | -34.1649 | -4.542
2.2036 | -34.1654 | -4.5553
2.2097 | -34.1659 | -4.5674
2.2158 | -34.167 | -4.578
2.222 | -34.1675 | -4.5888
2.2281 | -34.168 | -4.5998
2.2342 | -34.1685 | -4.6131
2.2403 | -34.1691 | -4.6239
2.2464 | -34.1699 | -4.6353
2.2525 | -34.1704 | -4.6462
2.2586 | -34.1713 | -4.6564
2.2647 | -34.1722 | -4.6661
2.2708 | -34.1728 | -4.678
2.2769 | -34.1737 | -4.6898
2.283 | -34.1743 | -4.6996
2.2891 | -34.1749 | -4.7105
2.2952 | -34.1756 | -4.721
2.3013 | -34.1764 | -4.7318
2.3074 | -34.1771 | -4.7421
2.3135 | -34.1777 | -4.7536
2.3196 | -34.1784 | -4.7628
2.3257 | -34.1788 | -4.7738
2.3318 | -34.1793 | -4.7832
2.3379 | -34.1799 | -4.7952
2.344 | -34.1803 | -4.8064
2.3501 | -34.1808 | -4.8177
2.3562 | -34.1812 | -4.8266
2.3623 | -34.1816 | -4.836
2.3685 | -34.182 | -4.8457
2.3746 | -34.1823 | -4.8566
2.3807 | -34.1829 | -4.8659
2.3868 | -34.1832 | -4.8754
2.3929 | -34.1837 | -4.8851
2.399 | -34.184 | -4.8938
2.4051 | -34.1846 | -4.9037
2.4112 | -34.185 | -4.9132
2.4173 | -34.1854 | -4.9227
2.4234 | -34.1855 | -4.9316
2.4295 | -34.1859 | -4.9406
2.4356 | -34.1861 | -4.9503
2.4417 | -34.1862 | -4.9606
2.4478 | -34.1864 | -4.9707
2.4539 | -34.1868 | -4.979
2.46 | -34.1866 | -4.989
2.4661 | -34.1867 | -4.9985
2.4722 | -34.187 | -5.0078
2.4783 | -34.1872 | -5.0168
2.4844 | -34.1873 | -5.0259
2.4905 | -34.1874 | -5.0354
2.4966 | -34.1872 | -5.0447
2.5027 | -34.1873 | -5.0534
2.5089 | -34.1873 | -5.0621
2.515 | -34.1872 | -5.072
2.5211 | -34.1873 | -5.0807
2.5272 | -34.1871 | -5.0901
2.5333 | -34.187 | -5.0999
2.5394 | -34.187 | -5.1092
2.5455 | -34.187 | -5.1187
2.5516 | -34.187 | -5.128
2.5577 | -34.187 | -5.138
2.5638 | -34.1869 | -5.1484
2.5699 | -34.1868 | -5.1589
2.576 | -34.1867 | -5.1694
2.5821 | -34.1867 | -5.1794
2.5882 | -34.1866 | -5.1893
2.5943 | -34.1865 | -5.1994
2.6004 | -34.1864 | -5.2096
2.6065 | -34.1864 | -5.2186
2.6126 | -34.1864 | -5.2282
2.6187 | -34.1863 | -5.2374
2.6248 | -34.1863 | -5.247
2.6309 | -34.1862 | -5.2583
2.637 | -34.1864 | -5.2694
2.6431 | -34.1864 | -5.2794
2.6492 | -34.1864 | -5.2897
2.6554 | -34.1865 | -5.3003
2.6615 | -34.1864 | -5.3105
2.6676 | -34.1864 | -5.3209
2.6737 | -34.1866 | -5.3318
2.6798 | -34.1867 | -5.3418
2.6859 | -34.1868 | -5.3526
2.692 | -34.1868 | -5.3627
2.6981 | -34.1869 | -5.374
2.7042 | -34.1871 | -5.3862
2.7103 | -34.1871 | -5.3965
2.7164 | -34.1874 | -5.4073
2.7225 | -34.1876 | -5.4188
2.7286 | -34.1877 | -5.4302
2.7347 | -34.1879 | -5.4405
2.7408 | -34.1882 | -5.4502
2.7469 | -34.1883 | -5.4617
2.753 | -34.1886 | -5.4726
2.7591 | -34.1888 | -5.4847
2.7652 | -34.1889 | -5.4962
2.7713 | -34.1891 | -5.5065
2.7774 | -34.1896 | -5.518
2.7835 | -34.1898 | -5.529
2.7896 | -34.19 | -5.5387
2.7958 | -34.1905 | -5.5493
2.8019 | -34.1911 | -5.5595
2.808 | -34.1914 | -5.5716
2.8141 | -34.1919 | -5.5846
2.8202 | -34.1923 | -5.5971
2.8263 | -34.1929 | -5.6082
2.8324 | -34.1932 | -5.6207
2.8385 | -34.1936 | -5.633
2.8446 | -34.1943 | -5.6449
2.8507 | -34.1946 | -5.6572
2.8568 | -34.1952 | -5.6686
2.8629 | -34.1957 | -5.6804
2.869 | -34.1962 | -5.6917
2.8751 | -34.1967 | -5.7041
2.8812 | -34.1974 | -5.7174
2.8873 | -34.1977 | -5.729
2.8934 | -34.1982 | -5.7411
2.8995 | -34.1986 | -5.7533
2.9056 | -34.199 | -5.7642
2.9117 | -34.1995 | -5.776
2.9178 | -34.2002 | -5.7889
2.9239 | -34.2007 | -5.8022
2.93 | -34.2012 | -5.8156
2.9361 | -34.2018 | -5.8273
2.9423 | -34.2025 | -5.8395
2.9484 | -34.203 | -5.8531
2.9545 | -34.2037 | -5.8665
2.9606 | -34.2043 | -5.8799
2.9667 | -34.2048 | -5.8918
2.9728 | -34.2053 | -5.9042
2.9789 | -34.2057 | -5.917
2.985 | -34.2062 | -5.9309
2.9911 | -34.2067 | -5.9436
2.9972 | -34.2073 | -5.9564
3.0033 | -34.2079 | -5.9694
3.0094 | -34.2085 | -5.9825
3.0155 | -34.2091 | -5.9953
3.0216 | -34.2096 | -6.0089
3.0277 | -34.2101 | -6.0225
3.0338 | -34.2108 | -6.0362
3.0399 | -34.2114 | -6.05
3.046 | -34.2118 | -6.0631
3.0521 | -34.2124 | -6.0756
3.0582 | -34.2128 | -6.0893
3.0643 | -34.2133 | -6.1023
3.0704 | -34.2136 | -6.1156
3.0765 | -34.2141 | -6.128
3.0827 | -34.2145 | -6.1413
3.0888 | -34.2149 | -6.1544
3.0949 | -34.2152 | -6.1676
3.101 | -34.2157 | -6.1817
3.1071 | -34.2161 | -6.1958
3.1132 | -34.2166 | -6.2097
3.1193 | -34.217 | -6.2228
3.1254 | -34.2174 | -6.2351
3.1315 | -34.2179 | -6.2493
3.1376 | -34.2182 | -6.2623
3.1437 | -34.2186 | -6.2758
3.1498 | -34.219 | -6.2898
3.1559 | -34.2194 | -6.3032
3.162 | -34.2195 | -6.3171
3.1681 | -34.2196 | -6.3312
3.1742 | -34.22 | -6.345
3.1803 | -34.2205 | -6.3581
3.1864 | -34.2209 | -6.3705
3.1925 | -34.2213 | -6.3842
3.1986 | -34.2214 | -6.3968
3.2047 | -34.2216 | -6.4106
3.2108 | -34.2219 | -6.4225
3.2169 | -34.2224 | -6.4362
3.223 | -34.2224 | -6.4495
3.2292 | -34.2225 | -6.4634
3.2353 | -34.2225 | -6.4732
3.2414 | -34.2227 | -6.4862
3.2475 | -34.2228 | -6.4973
3.2536 | -34.2231 | -6.511
3.2597 | -34.2234 | -6.5224
3.2658 | -34.2236 | -6.5334
3.2719 | -34.2239 | -6.5446
3.278 | -34.224 | -6.5567
3.2841 | -34.2242 | -6.5688
3.2902 | -34.2244 | -6.5814
3.2963 | -34.2246 | -6.593
3.3024 | -34.2246 | -6.6033
3.3085 | -34.2249 | -6.6158
3.3146 | -34.2249 | -6.6269
3.3207 | -34.2252 | -6.6389
3.3268 | -34.2255 | -6.6511
3.3329 | -34.2256 | -6.6612
3.339 | -34.2258 | -6.671
3.3451 | -34.226 | -6.6824
3.3512 | -34.2262 | -6.6921
3.3573 | -34.2265 | -6.7036
3.3634 | -34.2266 | -6.7152
3.3696 | -34.2267 | -6.727
3.3757 | -34.2268 | -6.737
3.3818 | -34.2272 | -6.7489
3.3879 | -34.2272 | -6.7593
3.394 | -34.2273 | -6.7697
3.4001 | -34.2274 | -6.7808
3.4062 | -34.2276 | -6.7911
3.4123 | -34.2278 | -6.8011
3.4184 | -34.2279 | -6.8122
3.4245 | -34.2281 | -6.8229
3.4306 | -34.2282 | -6.8329
3.4367 | -34.2284 | -6.843
3.4428 | -34.2288 | -6.8524
3.4489 | -34.2289 | -6.8636
3.455 | -34.2289 | -6.8741
3.4611 | -34.229 | -6.8852
3.4672 | -34.2291 | -6.8966
3.4733 | -34.2293 | -6.9067
3.4794 | -34.2294 | -6.9172
3.4855 | -34.2298 | -6.9271
3.4916 | -34.23 | -6.9371
3.4977 | -34.23 | -6.9481
3.5038 | -34.2303 | -6.9581
3.5099 | -34.2305 | -6.9681
3.5161 | -34.2307 | -6.9778
3.5222 | -34.231 | -6.9881
3.5283 | -34.231 | -6.9983
3.5344 | -34.2312 | -7.0092
3.5405 | -34.2315 | -7.0193
3.5466 | -34.2317 | -7.0291
3.5527 | -34.2319 | -7.0398
3.5588 | -34.2321 | -7.0498
3.5649 | -34.2323 | -7.0601
3.571 | -34.2325 | -7.0706
3.5771 | -34.2328 | -7.0813
3.5832 | -34.233 | -7.0915
3.5893 | -34.2333 | -7.1015
3.5954 | -34.2335 | -7.1121
3.6015 | -34.2337 | -7.1225
3.6076 | -34.2341 | -7.1323
3.6137 | -34.2343 | -7.1426
3.6198 | -34.2347 | -7.1519
3.6259 | -34.2349 | -7.1627
3.632 | -34.2352 | -7.1734
3.6381 | -34.2355 | -7.1833
3.6442 | -34.2358 | -7.1939
3.6503 | -34.2361 | -7.2033
3.6565 | -34.2365 | -7.215
3.6626 | -34.2367 | -7.2259
3.6687 | -34.2371 | -7.2366
3.6748 | -34.2375 | -7.2467
3.6809 | -34.2379 | -7.2573
3.687 | -34.2382 | -7.2685
3.6931 | -34.2386 | -7.2788
3.6992 | -34.239 | -7.2896
3.7053 | -34.2393 | -7.2999
3.7114 | -34.2396 | -7.3112
3.7175 | -34.24 | -7.3213
3.7236 | -34.2403 | -7.3305
3.7297 | -34.2405 | -7.3413
3.7358 | -34.2408 | -7.353
3.7419 | -34.2411 | -7.3647
3.748 | -34.2413 | -7.3756
3.7541 | -34.2416 | -7.3867
3.7602 | -34.242 | -7.3977
3.7663 | -34.2423 | -7.408
3.7724 | -34.2427 | -7.4192
3.7785 | -34.2431 | -7.4298
3.7846 | -34.2434 | -7.4406
3.7907 | -34.2438 | -7.4517
3.7969 | -34.2441 | -7.4633
3.803 | -34.2445 | -7.4737
3.8091 | -34.2449 | -7.4843
3.8152 | -34.2452 | -7.4956
3.8213 | -34.2455 | -7.5072
3.8274 | -34.246 | -7.518
3.8335 | -34.2464 | -7.5297
3.8396 | -34.2469 | -7.5414
3.8457 | -34.2473 | -7.5517
3.8518 | -34.2478 | -7.5628
3.8579 | -34.2482 | -7.5735
3.864 | -34.2486 | -7.5843
3.8701 | -34.2491 | -7.5953
3.8762 | -34.2495 | -7.6068
3.8823 | -34.2498 | -7.6174
3.8884 | -34.2503 | -7.6279
3.8945 | -34.2508 | -7.6393
3.9006 | -34.2512 | -7.6508
3.9067 | -34.2517 | -7.6626
3.9128 | -34.2521 | -7.6741
3.9189 | -34.2526 | -7.6849
3.925 | -34.2529 | -7.6959
3.9311 | -34.2534 | -7.7062
3.9372 | -34.2538 | -7.7178
3.9434 | -34.2542 | -7.7283
3.9495 | -34.2546 | -7.7401
3.9556 | -34.2551 | -7.7515
3.9617 | -34.2555 | -7.7628
3.9678 | -34.2559 | -7.7738
3.9739 | -34.2564 | -7.7852
3.98 | -34.2569 | -7.7972
3.9861 | -34.2572 | -7.8078
3.9922 | -34.2575 | -7.8196
3.9983 | -34.258 | -7.83
4.0044 | -34.2583 | -7.8401
4.0105 | -34.2588 | -7.8512
4.0166 | -34.2593 | -7.8623
4.0227 | -34.2598 | -7.8728
4.0288 | -34.2601 | -7.8842
4.0349 | -34.2605 | -7.8943
4.041 | -34.2609 | -7.9055
4.0471 | -34.2612 | -7.9159
4.0532 | -34.2617 | -7.9274
4.0593 | -34.2621 | -7.9391
4.0654 | -34.2625 | -7.9501
4.0715 | -34.2628 | -7.9615
4.0776 | -34.2632 | -7.972
4.0838 | -34.2636 | -7.9833
4.0899 | -34.2641 | -7.995
4.096 | -34.2644 | -8.0055
4.1021 | -34.2648 | -8.0166
4.1082 | -34.2653 | -8.0272
4.1143 | -34.2657 | -8.0376
4.1204 | -34.2662 | -8.0486
4.1265 | -34.2666 | -8.0593
4.1326 | -34.2671 | -8.07
4.1387 | -34.2675 | -8.0814
4.1448 | -34.268 | -8.0926
4.1509 | -34.2683 | -8.1035
4.157 | -34.2686 | -8.1147
4.1631 | -34.269 | -8.1261
4.1692 | -34.2694 | -8.137
4.1753 | -34.2699 | -8.1481
4.1814 | -34.2702 | -8.1583
4.1875 | -34.2706 | -8.1698
4.1936 | -34.271 | -8.1814
4.1997 | -34.2714 | -8.1919
4.2058 | -34.2717 | -8.2036
4.2119 | -34.2721 | -8.2136
4.218 | -34.2724 | -8.2241
4.2241 | -34.2729 | -8.234
4.2303 | -34.2733 | -8.2456
4.2364 | -34.2738 | -8.2573
4.2425 | -34.2741 | -8.268
4.2486 | -34.2745 | -8.2793
4.2547 | -34.2748 | -8.289
4.2608 | -34.2751 | -8.2991
4.2669 | -34.2755 | -8.31
4.273 | -34.2759 | -8.3209
4.2791 | -34.2762 | -8.3319
4.2852 | -34.2765 | -8.3424
4.2913 | -34.2768 | -8.3521
4.2974 | -34.2771 | -8.3632
4.3035 | -34.2774 | -8.3741
4.3096 | -34.2776 | -8.3848
4.3157 | -34.278 | -8.3951
4.3218 | -34.2783 | -8.4054
4.3279 | -34.2787 | -8.4161
4.334 | -34.279 | -8.4263
4.3401 | -34.2793 | -8.4367
4.3462 | -34.2797 | -8.4477
4.3523 | -34.2801 | -8.4579
4.3584 | -34.2803 | -8.4669
4.3645 | -34.2806 | -8.4769
4.3707 | -34.281 | -8.4867
4.3768 | -34.2812 | -8.4974
4.3829 | -34.2815 | -8.5079
4.389 | -34.2819 | -8.518
4.3951 | -34.2822 | -8.5279
4.4012 | -34.2825 | -8.5379
4.4073 | -34.2828 | -8.5484
4.4134 | -34.2832 | -8.5589
4.4195 | -34.2835 | -8.5688
4.4256 | -34.2838 | -8.5789
4.4317 | -34.2842 | -8.5887
4.4378 | -34.2844 | -8.5988
4.4439 | -34.2847 | -8.6088
4.45 | -34.285 | -8.6191
4.4561 | -34.2852 | -8.6293
4.4622 | -34.2854 | -8.6393
4.4683 | -34.2857 | -8.6488
4.4744 | -34.286 | -8.6592
4.4805 | -34.2864 | -8.6688
4.4866 | -34.2867 | -8.6784
4.4927 | -34.287 | -8.6889
4.4988 | -34.2873 | -8.6987
4.5049 | -34.2876 | -8.7084
4.511 | -34.2879 | -8.7179
4.5172 | -34.2883 | -8.7281
4.5233 | -34.2886 | -8.7389
4.5294 | -34.2889 | -8.7487
4.5355 | -34.2892 | -8.7585
4.5416 | -34.2896 | -8.7679
4.5477 | -34.2898 | -8.7776
4.5538 | -34.2901 | -8.787
4.5599 | -34.2905 | -8.7969
4.566 | -34.2908 | -8.807
4.5721 | -34.2911 | -8.8165
4.5782 | -34.2915 | -8.8266
4.5843 | -34.2918 | -8.8365
4.5904 | -34.2922 | -8.8461
4.5965 | -34.2926 | -8.8553
4.6026 | -34.2929 | -8.8661
4.6087 | -34.2932 | -8.8756
4.6148 | -34.2935 | -8.8864
4.6209 | -34.294 | -8.8953
4.627 | -34.2944 | -8.9056
4.6331 | -34.2946 | -8.9149
4.6392 | -34.2949 | -8.9244
4.6453 | -34.2953 | -8.9342
4.6514 | -34.2956 | -8.944
4.6576 | -34.2957 | -8.9547
4.6637 | -34.2961 | -8.9629
4.6698 | -34.2964 | -8.9719
4.6759 | -34.2966 | -8.9819
4.682 | -34.2969 | -8.9914
4.6881 | -34.2972 | -9.0006
4.6942 | -34.2975 | -9.0102
4.7003 | -34.2979 | -9.02
4.7064 | -34.2982 | -9.0297
4.7125 | -34.2986 | -9.0393
4.7186 | -34.2988 | -9.0482
4.7247 | -34.2992 | -9.0581
4.7308 | -34.2996 | -9.0677
4.7369 | -34.2999 | -9.0776
4.743 | -34.3002 | -9.0881
4.7491 | -34.3004 | -9.0979
4.7552 | -34.3007 | -9.1074
4.7613 | -34.301 | -9.1173
4.7674 | -34.3013 | -9.1263
4.7735 | -34.3015 | -9.1365
4.7796 | -34.3018 | -9.1464
4.7857 | -34.302 | -9.1563
4.7918 | -34.3024 | -9.1653
4.7979 | -34.3028 | -9.1747
4.8041 | -34.3029 | -9.1843
4.8102 | -34.3034 | -9.1938
4.8163 | -34.3037 | -9.2036
4.8224 | -34.3038 | -9.2125
4.8285 | -34.3043 | -9.222
4.8346 | -34.3046 | -9.2319
4.8407 | -34.3049 | -9.2415
4.8468 | -34.3052 | -9.2513
4.8529 | -34.3055 | -9.2612
4.859 | -34.3057 | -9.2708
4.8651 | -34.3061 | -9.2806
4.8712 | -34.3064 | -9.2903
4.8773 | -34.3067 | -9.3001
4.8834 | -34.307 | -9.3101
4.8895 | -34.3073 | -9.3206
4.8956 | -34.3076 | -9.3308
4.9017 | -34.3079 | -9.3406
4.9078 | -34.3082 | -9.3506
4.9139 | -34.3085 | -9.3596
4.92 | -34.3088 | -9.369
4.9261 | -34.3091 | -9.3783
4.9322 | -34.3096 | -9.3877
4.9383 | -34.31 | -9.397
4.9445 | -34.3103 | -9.4074
4.9506 | -34.3106 | -9.4166
4.9567 | -34.311 | -9.4262
4.9628 | -34.3113 | -9.4359
4.9689 | -34.3115 | -9.4458
4.975 | -34.3118 | -9.4552
4.9811 | -34.3122 | -9.4654
4.9872 | -34.3125 | -9.4755
4.9933 | -34.3129 | -9.4852
4.9994 | -34.3133 | -9.4951
5.0055 | -34.3137 | -9.505
5.0116 | -34.3141 | -9.5153
5.0177 | -34.3145 | -9.5254
5.0238 | -34.3148 | -9.536
5.0299 | -34.3153 | -9.547
5.036 | -34.3156 | -9.5568
5.0421 | -34.3159 | -9.5668
5.0482 | -34.3163 | -9.5773
5.0543 | -34.3166 | -9.587
5.0604 | -34.3171 | -9.5975
5.0665 | -34.3174 | -9.6075
5.0726 | -34.3178 | -9.6179
5.0787 | -34.3182 | -9.6283
5.0848 | -34.3185 | -9.6385
5.091 | -34.3189 | -9.6487
5.0971 | -34.3193 | -9.6592
5.1032 | -34.3197 | -9.6695
5.1093 | -34.32 | -9.6796
5.1154 | -34.3203 | -9.69
5.1215 | -34.3207 | -9.7003
5.1276 | -34.321 | -9.7104
5.1337 | -34.3215 | -9.7204
5.1398 | -34.3218 | -9.7306
5.1459 | -34.3222 | -9.7411
5.152 | -34.3226 | -9.7522
5.1581 | -34.3229 | -9.763
5.1642 | -34.3232 | -9.7732
5.1703 | -34.3237 | -9.7834
5.1764 | -34.324 | -9.7936
5.1825 | -34.3243 | -9.8043
5.1886 | -34.3247 | -9.8145
5.1947 | -34.325 | -9.8255
5.2008 | -34.3255 | -9.8355
5.2069 | -34.3259 | -9.8468
5.213 | -34.3263 | -9.8567
5.2191 | -34.3267 | -9.8672
5.2252 | -34.327 | -9.8777
5.2314 | -34.3273 | -9.8885
5.2375 | -34.3278 | -9.8988
5.2436 | -34.3281 | -9.9092
5.2497 | -34.3285 | -9.9194
5.2558 | -34.3289 | -9.9304
5.2619 | -34.3293 | -9.9413
5.268 | -34.3297 | -9.9518
5.2741 | -34.3301 | -9.9632
5.2802 | -34.3305 | -9.974
5.2863 | -34.3309 | -9.9844
5.2924 | -34.3312 | -9.9957
5.2985 | -34.3317 | -10.0072
5.3046 | -34.3322 | -10.0186
5.3107 | -34.3326 | -10.0295
5.3168 | -34.3331 | -10.0401
5.3229 | -34.3335 | -10.0502
5.329 | -34.3339 | -10.0603
5.3351 | -34.3344 | -10.0708
5.3412 | -34.3348 | -10.0812
5.3473 | -34.3353 | -10.0922
5.3534 | -34.3357 | -10.1032
5.3595 | -34.3361 | -10.1141
5.3656 | -34.3365 | -10.1253
5.3717 | -34.3368 | -10.1358
5.3779 | -34.3372 | -10.1467
5.384 | -34.3376 | -10.1577
5.3901 | -34.338 | -10.1684
5.3962 | -34.3384 | -10.1795
5.4023 | -34.3388 | -10.1902
5.4084 | -34.3392 | -10.2009
5.4145 | -34.3396 | -10.2122
5.4206 | -34.3399 | -10.2236
5.4267 | -34.3404 | -10.2349
5.4328 | -34.3408 | -10.2452
5.4389 | -34.3412 | -10.2558
5.445 | -34.3416 | -10.2671
5.4511 | -34.342 | -10.2778
5.4572 | -34.3425 | -10.2891
5.4633 | -34.3428 | -10.2996
5.4694 | -34.3432 | -10.3103
5.4755 | -34.3436 | -10.3214
5.4816 | -34.344 | -10.3322
5.4877 | -34.3444 | -10.3431
5.4938 | -34.3448 | -10.3538
5.4999 | -34.3452 | -10.3648
5.506 | -34.3456 | -10.3757
5.5121 | -34.346 | -10.387
5.5183 | -34.3464 | -10.3974
5.5244 | -34.3467 | -10.4083
5.5305 | -34.3471 | -10.4194
5.5366 | -34.3475 | -10.4301
5.5427 | -34.3479 | -10.4415
5.5488 | -34.3483 | -10.4527
5.5549 | -34.3486 | -10.4635
5.561 | -34.349 | -10.4749
5.5671 | -34.3495 | -10.4855
5.5732 | -34.3498 | -10.4965
5.5793 | -34.3502 | -10.5081
5.5854 | -34.3507 | -10.5192
5.5915 | -34.351 | -10.5301
5.5976 | -34.3514 | -10.5413
5.6037 | -34.3518 | -10.5521
5.6098 | -34.3521 | -10.5629
5.6159 | -34.3525 | -10.5741
5.622 | -34.353 | -10.5846
5.6281 | -34.3534 | -10.5956
5.6342 | -34.3538 | -10.6065
5.6403 | -34.3541 | -10.6169
5.6464 | -34.3546 | -10.6277
5.6525 | -34.3549 | -10.6383
5.6586 | -34.3554 | -10.6486
5.6648 | -34.3557 | -10.66
5.6709 | -34.3561 | -10.6711
5.677 | -34.3565 | -10.6828
5.6831 | -34.3569 | -10.6932
5.6892 | -34.3573 | -10.704
5.6953 | -34.3577 | -10.7149
5.7014 | -34.3581 | -10.7263
5.7075 | -34.3584 | -10.7374
5.7136 | -34.3588 | -10.7482
5.7197 | -34.3592 | -10.759
5.7258 | -34.3596 | -10.7698
5.7319 | -34.36 | -10.7806
5.738 | -34.3604 | -10.7915
5.7441 | -34.3608 | -10.8024
5.7502 | -34.3612 | -10.813
5.7563 | -34.3616 | -10.8236
5.7624 | -34.3619 | -10.8351
5.7685 | -34.3622 | -10.8457
5.7746 | -34.3626 | -10.8563
5.7807 | -34.363 | -10.8668
5.7868 | -34.3634 | -10.8775
5.7929 | -34.3637 | -10.8883
5.799 | -34.3641 | -10.8991
5.8052 | -34.3644 | -10.9101
5.8113 | -34.3648 | -10.921
5.8174 | -34.3651 | -10.9321
5.8235 | -34.3655 | -10.9427
5.8296 | -34.3659 | -10.9541
5.8357 | -34.3662 | -10.965
5.8418 | -34.3666 | -10.9759
5.8479 | -34.3669 | -10.987
5.854 | -34.3672 | -10.9981
5.8601 | -34.3676 | -11.0087
5.8662 | -34.3679 | -11.0198
5.8723 | -34.3683 | -11.0301
5.8784 | -34.3687 | -11.0409
5.8845 | -34.3691 | -11.0519
5.8906 | -34.3694 | -11.0626
5.8967 | -34.3698 | -11.0731
5.9028 | -34.3702 | -11.084
5.9089 | -34.3705 | -11.0942
5.915 | -34.3708 | -11.1051
5.9211 | -34.3713 | -11.116
5.9272 | -34.3716 | -11.1268
5.9333 | -34.3719 | -11.137
5.9394 | -34.3724 | -11.1475
5.9455 | -34.3727 | -11.158
5.9517 | -34.3731 | -11.1688
5.9578 | -34.3735 | -11.1797
5.9639 | -34.3739 | -11.1903
5.97 | -34.3743 | -11.2011
5.9761 | -34.3746 | -11.2116
5.9822 | -34.375 | -11.2223
5.9883 | -34.3755 | -11.2328
5.9944 | -34.3758 | -11.2435
6.0005 | -34.3763 | -11.2543
6.0066 | -34.3766 | -11.2648
6.0127 | -34.3769 | -11.2748
6.0188 | -34.3773 | -11.2858
6.0249 | -34.3777 | -11.2964
6.031 | -34.3781 | -11.307
6.0371 | -34.3785 | -11.3179
6.0432 | -34.3789 | -11.3286
6.0493 | -34.3793 | -11.3391
6.0554 | -34.3798 | -11.3497
6.0615 | -34.3801 | -11.3609
6.0676 | -34.3805 | -11.3721
6.0737 | -34.3809 | -11.3828
6.0798 | -34.3812 | -11.3934
6.0859 | -34.3816 | -11.404
6.0921 | -34.382 | -11.4145
6.0982 | -34.3824 | -11.4254
6.1043 | -34.3829 | -11.4359
6.1104 | -34.3833 | -11.4469
6.1165 | -34.3837 | -11.457
6.1226 | -34.3841 | -11.4671
6.1287 | -34.3845 | -11.4778
6.1348 | -34.3849 | -11.4885
6.1409 | -34.3853 | -11.4991
6.147 | -34.3857 | -11.5096
6.1531 | -34.3861 | -11.5201
6.1592 | -34.3865 | -11.5308
6.1653 | -34.3869 | -11.5412
6.1714 | -34.3873 | -11.5518
6.1775 | -34.3877 | -11.5627
6.1836 | -34.388 | -11.5731
6.1897 | -34.3884 | -11.5836
6.1958 | -34.3888 | -11.5943
6.2019 | -34.3892 | -11.6051
6.208 | -34.3896 | -11.6161
6.2141 | -34.39 | -11.6269
6.2202 | -34.3904 | -11.6379
6.2263 | -34.3908 | -11.6484
6.2325 | -34.3913 | -11.6586
6.2386 | -34.3917 | -11.669
6.2447 | -34.3921 | -11.6797
6.2508 | -34.3924 | -11.6902
6.2569 | -34.3929 | -11.7011
6.263 | -34.3933 | -11.7114
6.2691 | -34.3937 | -11.7219
6.2752 | -34.3942 | -11.7324
6.2813 | -34.3946 | -11.743
6.2874 | -34.3949 | -11.7536
6.2935 | -34.3954 | -11.7639
6.2996 | -34.3958 | -11.7746
6.3057 | -34.3962 | -11.785
6.3118 | -34.3966 | -11.7955
6.3179 | -34.397 | -11.806
6.324 | -34.3975 | -11.8168
6.3301 | -34.3979 | -11.8271
6.3362 | -34.3983 | -11.8377
6.3423 | -34.3988 | -11.8482
6.3484 | -34.3993 | -11.8589
6.3545 | -34.3997 | -11.869
6.3606 | -34.4001 | -11.8795
6.3667 | -34.4006 | -11.8898
6.3728 | -34.401 | -11.9004
6.379 | -34.4014 | -11.9109
6.3851 | -34.4018 | -11.9216
6.3912 | -34.4023 | -11.9319
6.3973 | -34.4028 | -11.9421
6.4034 | -34.4032 | -11.9527
6.4095 | -34.4036 | -11.9634
6.4156 | -34.404 | -11.9736
6.4217 | -34.4044 | -11.9841
6.4278 | -34.4048 | -11.9945
6.4339 | -34.4053 | -12.0049
6.44 | -34.4057 | -12.0151
6.4461 | -34.4061 | -12.026
6.4522 | -34.4065 | -12.036
6.4583 | -34.4069 | -12.0465
6.4644 | -34.4073 | -12.0573
6.4705 | -34.4078 | -12.0681
6.4766 | -34.4082 | -12.079
6.4827 | -34.4086 | -12.0899
6.4888 | -34.409 | -12.1003
6.4949 | -34.4094 | -12.1105
6.501 | -34.4098 | -12.1213
6.5071 | -34.4103 | -12.1319
6.5132 | -34.4107 | -12.1423
6.5194 | -34.4112 | -12.1526
6.5255 | -34.4116 | -12.1628
6.5316 | -34.412 | -12.1734
6.5377 | -34.4124 | -12.1842
6.5438 | -34.4129 | -12.1946
6.5499 | -34.4133 | -12.2049
6.556 | -34.4137 | -12.2153
6.5621 | -34.4141 | -12.2259
6.5682 | -34.4145 | -12.2361
6.5743 | -34.4149 | -12.2473
6.5804 | -34.4153 | -12.2578
6.5865 | -34.4156 | -12.2685
6.5926 | -34.4161 | -12.279
6.5987 | -34.4166 | -12.2892
6.6048 | -34.4169 | -12.3003
6.6109 | -34.4174 | -12.3106
6.617 | -34.4178 | -12.3219
6.6231 | -34.4181 | -12.3327
6.6292 | -34.4187 | -12.343
6.6353 | -34.4191 | -12.3544
6.6414 | -34.4196 | -12.3652
6.6475 | -34.42 | -12.3758
6.6536 | -34.4204 | -12.386
6.6597 | -34.4209 | -12.3964
6.6659 | -34.4213 | -12.4067
6.672 | -34.4218 | -12.4173
6.6781 | -34.4222 | -12.4275
6.6842 | -34.4226 | -12.4383
6.6903 | -34.4231 | -12.4487
6.6964 | -34.4235 | -12.4593
6.7025 | -34.4239 | -12.4699
6.7086 | -34.4243 | -12.4802
6.7147 | -34.4247 | -12.4908
6.7208 | -34.4251 | -12.5017
6.7269 | -34.4255 | -12.5119
6.733 | -34.4259 | -12.5218
6.7391 | -34.4264 | -12.5325
6.7452 | -34.4269 | -12.5425
6.7513 | -34.4273 | -12.5527
6.7574 | -34.4278 | -12.5635
6.7635 | -34.4282 | -12.5736
6.7696 | -34.4287 | -12.5841
6.7757 | -34.4292 | -12.5949
6.7818 | -34.4296 | -12.6054
6.7879 | -34.43 | -12.616
6.794 | -34.4304 | -12.6264
6.8001 | -34.4309 | -12.6368
6.8063 | -34.4313 | -12.6474
6.8124 | -34.4317 | -12.6579
6.8185 | -34.4321 | -12.6682
6.8246 | -34.4326 | -12.6788
6.8307 | -34.433 | -12.6892
6.8368 | -34.4334 | -12.6998
6.8429 | -34.4339 | -12.7102
6.849 | -34.4343 | -12.7206
6.8551 | -34.4347 | -12.7311
6.8612 | -34.4352 | -12.7414
6.8673 | -34.4356 | -12.7517
6.8734 | -34.436 | -12.7628
6.8795 | -34.4365 | -12.7735
6.8856 | -34.4369 | -12.7834
6.8917 | -34.4374 | -12.7933
6.8978 | -34.4378 | -12.8038
6.9039 | -34.4382 | -12.8142
6.91 | -34.4386 | -12.8247
6.9161 | -34.4391 | -12.835
6.9222 | -34.4395 | -12.8455
6.9283 | -34.4398 | -12.8557
6.9344 | -34.4403 | -12.8659
6.9405 | -34.4407 | -12.8761
6.9466 | -34.4411 | -12.8866
6.9528 | -34.4415 | -12.897
6.9589 | -34.4419 | -12.9071
6.965 | -34.4424 | -12.9175
6.9711 | -34.4428 | -12.9278
6.9772 | -34.4432 | -12.9379
6.9833 | -34.4436 | -12.9485
6.9894 | -34.444 | -12.9592
6.9955 | -34.4444 | -12.9696
7.0016 | -34.4448 | -12.9804
7.0077 | -34.4453 | -12.9911
7.0138 | -34.4457 | -13.0018
7.0199 | -34.4461 | -13.0121
7.026 | -34.4464 | -13.0224
7.0321 | -34.4468 | -13.0329
7.0382 | -34.4472 | -13.0432
7.0443 | -34.4477 | -13.0538
7.0504 | -34.4481 | -13.0645
7.0565 | -34.4485 | -13.0751
7.0626 | -34.4489 | -13.0854
7.0687 | -34.4493 | -13.0956
7.0748 | -34.4497 | -13.1059
7.0809 | -34.4501 | -13.1164
7.087 | -34.4505 | -13.1268
7.0932 | -34.4509 | -13.1371
7.0993 | -34.4513 | -13.1476
7.1054 | -34.4517 | -13.1577
7.1115 | -34.452 | -13.1681
7.1176 | -34.4523 | -13.1784
7.1237 | -34.4527 | -13.1888
7.1298 | -34.453 | -13.1987
7.1359 | -34.4534 | -13.2088
7.142 | -34.4538 | -13.2193
7.1481 | -34.4542 | -13.2294
7.1542 | -34.4546 | -13.2396
7.1603 | -34.455 | -13.2498
7.1664 | -34.4554 | -13.2602
7.1725 | -34.4558 | -13.2702
7.1786 | -34.4562 | -13.2806
7.1847 | -34.4566 | -13.2908
7.1908 | -34.457 | -13.3012
7.1969 | -34.4574 | -13.3114
7.203 | -34.4578 | -13.3218
7.2091 | -34.4582 | -13.3319
7.2152 | -34.4585 | -13.3422
7.2213 | -34.4589 | -13.3524
7.2274 | -34.4593 | -13.3626
7.2335 | -34.4597 | -13.3729
7.2397 | -34.4601 | -13.3832
7.2458 | -34.4604 | -13.3931
7.2519 | -34.4608 | -13.4034
7.258 | -34.4612 | -13.4137
7.2641 | -34.4616 | -13.4239
7.2702 | -34.462 | -13.4343
7.2763 | -34.4624 | -13.4443
7.2824 | -34.4629 | -13.4544
7.2885 | -34.4633 | -13.4645
7.2946 | -34.4637 | -13.475
7.3007 | -34.464 | -13.4852
7.3068 | -34.4644 | -13.4956
7.3129 | -34.4648 | -13.5055
7.319 | -34.4652 | -13.5158
7.3251 | -34.4656 | -13.5258
7.3312 | -34.466 | -13.5357
7.3373 | -34.4664 | -13.5456
7.3434 | -34.4669 | -13.5556
7.3495 | -34.4672 | -13.5654
7.3556 | -34.4676 | -13.5759
7.3617 | -34.468 | -13.5864
7.3678 | -34.4684 | -13.5967
7.3739 | -34.4688 | -13.6064
7.3801 | -34.4691 | -13.6163
7.3862 | -34.4695 | -13.6264
7.3923 | -34.4699 | -13.6364
7.3984 | -34.4703 | -13.6466
7.4045 | -34.4707 | -13.6567
7.4106 | -34.4711 | -13.6666
7.4167 | -34.4714 | -13.677
7.4228 | -34.4719 | -13.6869
7.4289 | -34.4722 | -13.6968
7.435 | -34.4727 | -13.7067
7.4411 | -34.4731 | -13.7165
7.4472 | -34.4735 | -13.7267
7.4533 | -34.4739 | -13.7369
7.4594 | -34.4743 | -13.7472
7.4655 | -34.4747 | -13.7574
7.4716 | -34.475 | -13.7676
7.4777 | -34.4754 | -13.7772
7.4838 | -34.4759 | -13.7876
7.4899 | -34.4763 | -13.7977
7.496 | -34.4767 | -13.8079
7.5021 | -34.4771 | -13.8179
7.5082 | -34.4774 | -13.8281
7.5143 | -34.4779 | -13.838
7.5204 | -34.4782 | -13.8477
7.5266 | -34.4786 | -13.8579
7.5327 | -34.479 | -13.8679
7.5388 | -34.4794 | -13.8777
7.5449 | -34.4797 | -13.8879
7.551 | -34.4801 | -13.8982
7.5571 | -34.4805 | -13.9081
7.5632 | -34.4809 | -13.918
7.5693 | -34.4813 | -13.9283
7.5754 | -34.4817 | -13.9384
7.5815 | -34.4821 | -13.9481
7.5876 | -34.4824 | -13.9579
7.5937 | -34.4828 | -13.9678
7.5998 | -34.4832 | -13.9779
7.6059 | -34.4835 | -13.9878
7.612 | -34.4839 | -13.9977
7.6181 | -34.4843 | -14.0077
7.6242 | -34.4847 | -14.0177
7.6303 | -34.4852 | -14.0274
7.6364 | -34.4855 | -14.0376
7.6425 | -34.486 | -14.0479
7.6486 | -34.4863 | -14.0579
7.6547 | -34.4867 | -14.0681
7.6608 | -34.4872 | -14.0781
7.667 | -34.4875 | -14.0875
7.6731 | -34.4879 | -14.0977
7.6792 | -34.4882 | -14.1077
7.6853 | -34.4886 | -14.1175
7.6914 | -34.4889 | -14.1275
7.6975 | -34.4893 | -14.1377
7.7036 | -34.4897 | -14.1478
7.7097 | -34.4901 | -14.1575
7.7158 | -34.4904 | -14.1674
7.7219 | -34.4908 | -14.1774
7.728 | -34.4911 | -14.1874
7.7341 | -34.4915 | -14.1974
7.7402 | -34.4919 | -14.2073
7.7463 | -34.4923 | -14.2173
7.7524 | -34.4926 | -14.2275
7.7585 | -34.493 | -14.2376
7.7646 | -34.4934 | -14.2473
7.7707 | -34.4937 | -14.257
7.7768 | -34.494 | -14.2667
7.7829 | -34.4943 | -14.2767
7.789 | -34.4947 | -14.2863
7.7951 | -34.4951 | -14.2966
7.8012 | -34.4955 | -14.3069
7.8073 | -34.4958 | -14.3166
7.8135 | -34.4962 | -14.3267
7.8196 | -34.4965 | -14.3365
7.8257 | -34.497 | -14.3467
7.8318 | -34.4973 | -14.3571
7.8379 | -34.4977 | -14.3666
7.844 | -34.498 | -14.3761
7.8501 | -34.4983 | -14.3862
7.8562 | -34.4986 | -14.3967
7.8623 | -34.4991 | -14.4067
7.8684 | -34.4994 | -14.4163
7.8745 | -34.4998 | -14.4261
7.8806 | -34.5001 | -14.4355
7.8867 | -34.5005 | -14.4454
7.8928 | -34.5009 | -14.4553
7.8989 | -34.5013 | -14.4647
7.905 | -34.5017 | -14.475
7.9111 | -34.5021 | -14.4849
7.9172 | -34.5024 | -14.4946
7.9233 | -34.5027 | -14.5046
7.9294 | -34.5032 | -14.5146
7.9355 | -34.5035 | -14.5241
7.9416 | -34.5038 | -14.5339
7.9477 | -34.5042 | -14.5434
7.9539 | -34.5046 | -14.553
7.96 | -34.505 | -14.5631
7.9661 | -34.5053 | -14.5731
7.9722 | -34.5057 | -14.5831
7.9783 | -34.5061 | -14.5929
7.9844 | -34.5065 | -14.6032
7.9905 | -34.5068 | -14.6129
7.9966 | -34.5071 | -14.6227
8.0027 | -34.5075 | -14.6329
8.0088 | -34.5079 | -14.6424
8.0149 | -34.5082 | -14.652
8.021 | -34.5086 | -14.6618
8.0271 | -34.509 | -14.672
8.0332 | -34.5093 | -14.6819
8.0393 | -34.5097 | -14.6919
8.0454 | -34.5101 | -14.7018
8.0515 | -34.5104 | -14.7115
8.0576 | -34.5107 | -14.7216
8.0637 | -34.5111 | -14.731
8.0698 | -34.5115 | -14.7406
8.0759 | -34.5119 | -14.7502
8.082 | -34.5122 | -14.7603
8.0881 | -34.5126 | -14.77
8.0942 | -34.5129 | -14.7802
8.1004 | -34.5133 | -14.7903
8.1065 | -34.5137 | -14.8002
8.1126 | -34.5141 | -14.8099
8.1187 | -34.5145 | -14.8199
8.1248 | -34.5148 | -14.83
8.1309 | -34.5152 | -14.8399
8.137 | -34.5155 | -14.8497
8.1431 | -34.5159 | -14.8597
8.1492 | -34.5163 | -14.8697
8.1553 | -34.5167 | -14.8795
8.1614 | -34.517 | -14.8895
8.1675 | -34.5174 | -14.8992
8.1736 | -34.5177 | -14.9094
8.1797 | -34.5181 | -14.9193
8.1858 | -34.5184 | -14.9291
8.1919 | -34.5188 | -14.9389
8.198 | -34.5191 | -14.9485
8.2041 | -34.5195 | -14.9587
8.2102 | -34.5199 | -14.9685
8.2163 | -34.5203 | -14.9784
8.2224 | -34.5207 | -14.9885
8.2285 | -34.521 | -14.9982
8.2346 | -34.5214 | -15.0081
8.2408 | -34.5218 | -15.0181
8.2469 | -34.5221 | -15.0278
8.253 | -34.5225 | -15.0376
8.2591 | -34.5229 | -15.0471
8.2652 | -34.5233 | -15.0567
8.2713 | -34.5237 | -15.0666
8.2774 | -34.5241 | -15.0765
8.2835 | -34.5245 | -15.0866
8.2896 | -34.5248 | -15.096
8.2957 | -34.5252 | -15.1058
8.3018 | -34.5256 | -15.1156
8.3079 | -34.526 | -15.1255
8.314 | -34.5263 | -15.1352
8.3201 | -34.5267 | -15.1449
8.3262 | -34.5271 | -15.1547
8.3323 | -34.5275 | -15.1646
8.3384 | -34.5278 | -15.1748
8.3445 | -34.5282 | -15.1842
8.3506 | -34.5286 | -15.1942
8.3567 | -34.5289 | -15.2043
8.3628 | -34.5293 | -15.2141
8.3689 | -34.5297 | -15.224
8.375 | -34.53 | -15.2339
8.3812 | -34.5304 | -15.2437
8.3873 | -34.5308 | -15.2537
8.3934 | -34.5312 | -15.2635
8.3995 | -34.5315 | -15.2735
8.4056 | -34.5319 | -15.2834
8.4117 | -34.5323 | -15.2934
8.4178 | -34.5327 | -15.3029
8.4239 | -34.533 | -15.3125
8.43 | -34.5335 | -15.3223
8.4361 | -34.5339 | -15.3325
8.4422 | -34.5342 | -15.3423
8.4483 | -34.5345 | -15.3522
8.4544 | -34.5349 | -15.3623
8.4605 | -34.5353 | -15.3721
8.4666 | -34.5356 | -15.382
8.4727 | -34.536 | -15.3918
8.4788 | -34.5364 | -15.4015
8.4849 | -34.5367 | -15.4116
8.491 | -34.5371 | -15.4213
8.4971 | -34.5375 | -15.4313
8.5032 | -34.5378 | -15.4413
8.5093 | -34.5382 | -15.4511
8.5154 | -34.5386 | -15.4607
8.5215 | -34.5389 | -15.4707
8.5277 | -34.5393 | -15.4806
8.5338 | -34.5397 | -15.4902
8.5399 | -34.54 | -15.5003
8.546 | -34.5404 | -15.5099
8.5521 | -34.5408 | -15.5199
8.5582 | -34.5412 | -15.5299
8.5643 | -34.5415 | -15.5396
8.5704 | -34.5419 | -15.5493
8.5765 | -34.5423 | -15.559
8.5826 | -34.5427 | -15.5689
8.5887 | -34.5431 | -15.5786
8.5948 | -34.5434 | -15.5881
8.6009 | -34.5438 | -15.5979
8.607 | -34.5441 | -15.6079
8.6131 | -34.5445 | -15.618
8.6192 | -34.5449 | -15.6275
8.6253 | -34.5453 | -15.6376
8.6314 | -34.5457 | -15.6474
8.6375 | -34.546 | -15.6575
8.6436 | -34.5465 | -15.6673
8.6497 | -34.5469 | -15.6772
8.6558 | -34.5472 | -15.6869
8.6619 | -34.5476 | -15.6969
8.6681 | -34.5479 | -15.7066
8.6742 | -34.5483 | -15.7167
8.6803 | -34.5487 | -15.7268
8.6864 | -34.5491 | -15.737
8.6925 | -34.5495 | -15.747
8.6986 | -34.5499 | -15.7571
8.7047 | -34.5503 | -15.767
8.7108 | -34.5507 | -15.777
8.7169 | -34.551 | -15.7865
8.723 | -34.5515 | -15.7965
8.7291 | -34.5518 | -15.8063
8.7352 | -34.5523 | -15.8163
8.7413 | -34.5526 | -15.8261
8.7474 | -34.553 | -15.8357
8.7535 | -34.5535 | -15.8457
8.7596 | -34.5539 | -15.8561
8.7657 | -34.5542 | -15.8659
8.7718 | -34.5546 | -15.8754
8.7779 | -34.555 | -15.8858
8.784 | -34.5553 | -15.8955
8.7901 | -34.5557 | -15.9057
8.7962 | -34.5561 | -15.9158
8.8023 | -34.5565 | -15.9254
8.8084 | -34.5569 | -15.9355
8.8146 | -34.5573 | -15.9453
8.8207 | -34.5577 | -15.9552
8.8268 | -34.5581 | -15.965
8.8329 | -34.5585 | -15.975
8.839 | -34.5589 | -15.9849
8.8451 | -34.5593 | -15.9947
8.8512 | -34.5597 | -16.0047
8.8573 | -34.5601 | -16.0144
8.8634 | -34.5606 | -16.0242
8.8695 | -34.5609 | -16.0342
8.8756 | -34.5613 | -16.0438
8.8817 | -34.5617 | -16.054
8.8878 | -34.562 | -16.064
8.8939 | -34.5624 | -16.0736
8.9 | -34.5628 | -16.0832
8.9061 | -34.5632 | -16.0931
8.9122 | -34.5636 | -16.1029
8.9183 | -34.564 | -16.1126
8.9244 | -34.5644 | -16.1221
8.9305 | -34.5648 | -16.1321
8.9366 | -34.5652 | -16.1418
8.9427 | -34.5656 | -16.1517
8.9488 | -34.566 | -16.1615
8.955 | -34.5664 | -16.1714
8.9611 | -34.5668 | -16.1813
8.9672 | -34.5672 | -16.1915
8.9733 | -34.5676 | -16.2013
8.9794 | -34.568 | -16.2114
8.9855 | -34.5684 | -16.2211
8.9916 | -34.5688 | -16.2307
8.9977 | -34.5692 | -16.2406
9.0038 | -34.5697 | -16.2507
9.0099 | -34.5701 | -16.2606
9.016 | -34.5705 | -16.2704
9.0221 | -34.5709 | -16.2807
9.0282 | -34.5713 | -16.2908
9.0343 | -34.5717 | -16.3009
9.0404 | -34.5722 | -16.3106
9.0465 | -34.5726 | -16.3207
9.0526 | -34.573 | -16.3305
9.0587 | -34.5734 | -16.3403
9.0648 | -34.5739 | -16.3502
9.0709 | -34.5743 | -16.3599
9.077 | -34.5747 | -16.3702
9.0831 | -34.5751 | -16.3801
9.0892 | -34.5755 | -16.3899
9.0953 | -34.5758 | -16.3998
9.1015 | -34.5762 | -16.4097
9.1076 | -34.5767 | -16.4197
9.1137 | -34.5771 | -16.4295
9.1198 | -34.5775 | -16.439
9.1259 | -34.578 | -16.4489
9.132 | -34.5784 | -16.459
9.1381 | -34.5788 | -16.4688
9.1442 | -34.5793 | -16.4784
9.1503 | -34.5797 | -16.4883
9.1564 | -34.5801 | -16.4983
9.1625 | -34.5805 | -16.5081
9.1686 | -34.5809 | -16.518
9.1747 | -34.5813 | -16.528
9.1808 | -34.5818 | -16.538
9.1869 | -34.5822 | -16.5481
9.193 | -34.5826 | -16.5582
9.1991 | -34.5829 | -16.5677
9.2052 | -34.5834 | -16.5775
9.2113 | -34.5838 | -16.5877
9.2174 | -34.5842 | -16.5981
9.2235 | -34.5847 | -16.6077
9.2296 | -34.5851 | -16.6178
9.2357 | -34.5855 | -16.6276
9.2419 | -34.586 | -16.6375
9.248 | -34.5864 | -16.6475
9.2541 | -34.5868 | -16.6575
9.2602 | -34.5872 | -16.6675
9.2663 | -34.5876 | -16.6773
9.2724 | -34.588 | -16.6872
9.2785 | -34.5885 | -16.6972
9.2846 | -34.5889 | -16.7071
9.2907 | -34.5893 | -16.717
9.2968 | -34.5897 | -16.7268
9.3029 | -34.5901 | -16.7368
9.309 | -34.5905 | -16.7466
9.3151 | -34.591 | -16.7565
9.3212 | -34.5914 | -16.7665
9.3273 | -34.5919 | -16.7763
9.3334 | -34.5924 | -16.786
9.3395 | -34.5928 | -16.7959
9.3456 | -34.5933 | -16.8057
9.3517 | -34.5937 | -16.8151
9.3578 | -34.5942 | -16.8248
9.3639 | -34.5946 | -16.835
9.37 | -34.5951 | -16.8449
9.3761 | -34.5955 | -16.8546
9.3822 | -34.5959 | -16.8643
9.3884 | -34.5964 | -16.8742
9.3945 | -34.5968 | -16.8841
9.4006 | -34.5972 | -16.894
9.4067 | -34.5976 | -16.904
9.4128 | -34.5981 | -16.9138
9.4189 | -34.5985 | -16.9237
9.425 | -34.599 | -16.9334
9.4311 | -34.5994 | -16.9433
9.4372 | -34.5998 | -16.9531
9.4433 | -34.6002 | -16.9632
9.4494 | -34.6006 | -16.9731
9.4555 | -34.6011 | -16.9829
9.4616 | -34.6016 | -16.9927
9.4677 | -34.602 | -17.0024
9.4738 | -34.6024 | -17.0126
9.4799 | -34.6028 | -17.0224
9.486 | -34.6032 | -17.0326
9.4921 | -34.6037 | -17.0429
9.4982 | -34.6041 | -17.0524
9.5043 | -34.6045 | -17.062
9.5104 | -34.605 | -17.0717
9.5165 | -34.6054 | -17.0815
9.5226 | -34.6059 | -17.0914
9.5288 | -34.6063 | -17.1013
9.5349 | -34.6068 | -17.1111
9.541 | -34.6072 | -17.121
9.5471 | -34.6077 | -17.131
9.5532 | -34.6081 | -17.1408
9.5593 | -34.6086 | -17.1506
9.5654 | -34.609 | -17.1604
9.5715 | -34.6095 | -17.1705
9.5776 | -34.6099 | -17.1806
9.5837 | -34.6103 | -17.1902
9.5898 | -34.6107 | -17.2002
9.5959 | -34.6112 | -17.2099
9.602 | -34.6116 | -17.2198
9.6081 | -34.612 | -17.2298
9.6142 | -34.6125 | -17.2395
9.6203 | -34.6129 | -17.2494
9.6264 | -34.6133 | -17.259
9.6325 | -34.6137 | -17.2689
9.6386 | -34.6142 | -17.2787
9.6447 | -34.6146 | -17.2884
9.6508 | -34.615 | -17.2984
9.6569 | -34.6155 | -17.3082
9.663 | -34.6159 | -17.318
9.6691 | -34.6164 | -17.3278
9.6753 | -34.6169 | -17.3376
9.6814 | -34.6173 | -17.3474
9.6875 | -34.6178 | -17.3571
9.6936 | -34.6183 | -17.3669
9.6997 | -34.6187 | -17.3768
9.7058 | -34.6192 | -17.3866
9.7119 | -34.6196 | -17.3962
9.718 | -34.62 | -17.4063
9.7241 | -34.6204 | -17.4162
9.7302 | -34.6208 | -17.4261
9.7363 | -34.6213 | -17.4357
9.7424 | -34.6218 | -17.4455
9.7485 | -34.6222 | -17.4552
9.7546 | -34.6227 | -17.4649
9.7607 | -34.6231 | -17.4748
9.7668 | -34.6235 | -17.485
9.7729 | -34.624 | -17.4946
9.779 | -34.6245 | -17.5044
9.7851 | -34.6249 | -17.5143
9.7912 | -34.6254 | -17.5247
9.7973 | -34.6257 | -17.5346
9.8034 | -34.6262 | -17.5447
9.8095 | -34.6266 | -17.5543
9.8157 | -34.6271 | -17.564
9.8218 | -34.6275 | -17.5734
9.8279 | -34.628 | -17.5834
9.834 | -34.6285 | -17.593
9.8401 | -34.6289 | -17.6027
9.8462 | -34.6294 | -17.6123
9.8523 | -34.6298 | -17.6221
9.8584 | -34.6303 | -17.6321
9.8645 | -34.6307 | -17.6418
9.8706 | -34.6311 | -17.6515
9.8767 | -34.6316 | -17.661
9.8828 | -34.6321 | -17.6704
9.8889 | -34.6326 | -17.6805
9.895 | -34.633 | -17.6899
9.9011 | -34.6334 | -17.6992
9.9072 | -34.6339 | -17.7086
9.9133 | -34.6343 | -17.7183
9.9194 | -34.6348 | -17.7278
9.9255 | -34.6352 | -17.7377
9.9316 | -34.6357 | -17.7472
9.9377 | -34.6362 | -17.7568
9.9438 | -34.6367 | -17.7665
9.9499 | -34.6371 | -17.7762
9.956 | -34.6376 | -17.7856
9.9622 | -34.638 | -17.7952
9.9683 | -34.6385 | -17.805
9.9744 | -34.6389 | -17.8145
9.9805 | -34.6394 | -17.8243
9.9866 | -34.6398 | -17.8336
9.9927 | -34.6403 | -17.8436
9.9988 | -34.6408 | -17.853
10.0049 | -34.6412 | -17.8626
10.011 | -34.6417 | -17.8722
10.0171 | -34.6421 | -17.8819
10.0232 | -34.6426 | -17.8914
10.0293 | -34.6431 | -17.9007
10.0354 | -34.6436 | -17.91
10.0415 | -34.644 | -17.9196
10.0476 | -34.6445 | -17.9296
10.0537 | -34.6449 | -17.9394
10.0598 | -34.6454 | -17.9492
10.0659 | -34.6458 | -17.9589
10.072 | -34.6463 | -17.9684
10.0781 | -34.6468 | -17.9781
10.0842 | -34.6472 | -17.9878
10.0903 | -34.6477 | -17.9977
10.0964 | -34.6482 | -18.0071
10.1026 | -34.6487 | -18.0168
10.1087 | -34.6492 | -18.0265
10.1148 | -34.6497 | -18.0362
10.1209 | -34.6501 | -18.0457
10.127 | -34.6506 | -18.0551
10.1331 | -34.6511 | -18.0646
10.1392 | -34.6515 | -18.0742
10.1453 | -34.652 | -18.0839
10.1514 | -34.6524 | -18.0932
10.1575 | -34.6529 | -18.1026
10.1636 | -34.6534 | -18.1116
10.1697 | -34.6538 | -18.1211
10.1758 | -34.6543 | -18.1307
10.1819 | -34.6548 | -18.1404
10.188 | -34.6553 | -18.1497
10.1941 | -34.6557 | -18.1594
10.2002 | -34.6561 | -18.1683
10.2063 | -34.6566 | -18.1781
10.2124 | -34.657 | -18.1877
10.2185 | -34.6575 | -18.1971
10.2246 | -34.6579 | -18.2067
10.2307 | -34.6584 | -18.2164
10.2368 | -34.6589 | -18.2252
10.2429 | -34.6594 | -18.2345
10.2491 | -34.6598 | -18.2438
10.2552 | -34.6603 | -18.2537
10.2613 | -34.6608 | -18.2634
10.2674 | -34.6612 | -18.273
10.2735 | -34.6616 | -18.2827
10.2796 | -34.6621 | -18.2923
10.2857 | -34.6625 | -18.3018
10.2918 | -34.663 | -18.3113
10.2979 | -34.6635 | -18.321
10.304 | -34.664 | -18.3307
10.3101 | -34.6645 | -18.3402
10.3162 | -34.665 | -18.3496
10.3223 | -34.6654 | -18.3592
10.3284 | -34.6659 | -18.3687
10.3345 | -34.6663 | -18.3781
10.3406 | -34.6668 | -18.3874
10.3467 | -34.6673 | -18.397
10.3528 | -34.6677 | -18.4063
10.3589 | -34.6681 | -18.416
10.365 | -34.6686 | -18.4257
10.3711 | -34.669 | -18.4353
10.3772 | -34.6694 | -18.4448
10.3833 | -34.6699 | -18.4542
10.3895 | -34.6703 | -18.4636
10.3956 | -34.6708 | -18.4728
10.4017 | -34.6712 | -18.4823
10.4078 | -34.6717 | -18.4919
10.4139 | -34.6722 | -18.5012
10.42 | -34.6727 | -18.5107
10.4261 | -34.6732 | -18.52
10.4322 | -34.6737 | -18.5294
10.4383 | -34.6741 | -18.5389
10.4444 | -34.6746 | -18.548
10.4505 | -34.6751 | -18.5575
10.4566 | -34.6755 | -18.5669
10.4627 | -34.676 | -18.5764
10.4688 | -34.6765 | -18.5854
10.4749 | -34.677 | -18.5947
10.481 | -34.6774 | -18.604
10.4871 | -34.6779 | -18.6136
10.4932 | -34.6784 | -18.6231
10.4993 | -34.6789 | -18.6324
10.5054 | -34.6794 | -18.6418
10.5115 | -34.6798 | -18.6512
10.5176 | -34.6803 | -18.6604
10.5237 | -34.6808 | -18.6696
10.5298 | -34.6812 | -18.6788
10.536 | -34.6817 | -18.6884
10.5421 | -34.6822 | -18.6974
10.5482 | -34.6826 | -18.7071
10.5543 | -34.6832 | -18.7165
10.5604 | -34.6836 | -18.7256
10.5665 | -34.6841 | -18.7349
10.5726 | -34.6845 | -18.7441
10.5787 | -34.685 | -18.7537
10.5848 | -34.6855 | -18.7628
10.5909 | -34.686 | -18.7722
10.597 | -34.6864 | -18.7815
10.6031 | -34.6869 | -18.7909
10.6092 | -34.6874 | -18.8006
10.6153 | -34.6878 | -18.8102
10.6214 | -34.6884 | -18.8194
10.6275 | -34.6889 | -18.8286
10.6336 | -34.6893 | -18.8384
10.6397 | -34.6898 | -18.8476
10.6458 | -34.6902 | -18.8568
10.6519 | -34.6906 | -18.8661
10.658 | -34.6911 | -18.8751
10.6641 | -34.6916 | -18.8844
10.6702 | -34.6921 | -18.894
10.6764 | -34.6926 | -18.9032
10.6825 | -34.6931 | -18.9127
10.6886 | -34.6936 | -18.9223
10.6947 | -34.694 | -18.9313
10.7008 | -34.6945 | -18.9406
10.7069 | -34.695 | -18.9499
10.713 | -34.6955 | -18.9592
10.7191 | -34.696 | -18.9687
10.7252 | -34.6965 | -18.9778
10.7313 | -34.697 | -18.9872
10.7374 | -34.6975 | -18.9962
10.7435 | -34.6979 | -19.0058
10.7496 | -34.6984 | -19.015
10.7557 | -34.6989 | -19.0243
10.7618 | -34.6994 | -19.0336
10.7679 | -34.6998 | -19.0429
10.774 | -34.7003 | -19.0522
10.7801 | -34.7007 | -19.0619
10.7862 | -34.7013 | -19.0712
10.7923 | -34.7018 | -19.0805
10.7984 | -34.7022 | -19.0897
10.8045 | -34.7027 | -19.0992
10.8106 | -34.7032 | -19.1085
10.8168 | -34.7037 | -19.1178
10.8229 | -34.7041 | -19.1272
10.829 | -34.7046 | -19.1366
10.8351 | -34.7051 | -19.1459
10.8412 | -34.7055 | -19.1551
10.8473 | -34.7059 | -19.1642
10.8534 | -34.7064 | -19.1737
10.8595 | -34.7069 | -19.1828
10.8656 | -34.7074 | -19.192
10.8717 | -34.7079 | -19.2009
10.8778 | -34.7084 | -19.21
10.8839 | -34.7089 | -19.2192
10.89 | -34.7093 | -19.2282
10.8961 | -34.7098 | -19.2375
10.9022 | -34.7103 | -19.2467
10.9083 | -34.7108 | -19.2562
10.9144 | -34.7113 | -19.2654
10.9205 | -34.7118 | -19.2747
10.9266 | -34.7123 | -19.2843
10.9327 | -34.7129 | -19.2936
10.9388 | -34.7133 | -19.3029
10.9449 | -34.7138 | -19.3122
10.951 | -34.7143 | -19.3214
10.9571 | -34.7147 | -19.3307
10.9633 | -34.7152 | -19.3402
10.9694 | -34.7157 | -19.3497
10.9755 | -34.7162 | -19.3589
10.9816 | -34.7167 | -19.3682
10.9877 | -34.7172 | -19.3775
10.9938 | -34.7177 | -19.3868
10.9999 | -34.7182 | -19.396
11.006 | -34.7187 | -19.4052
11.0121 | -34.7191 | -19.4148
11.0182 | -34.7196 | -19.4237
11.0243 | -34.7201 | -19.433
11.0304 | -34.7206 | -19.4421
11.0365 | -34.7211 | -19.4513
11.0426 | -34.7216 | -19.4604
11.0487 | -34.7221 | -19.4696
11.0548 | -34.7225 | -19.4788
11.0609 | -34.723 | -19.4883
11.067 | -34.7235 | -19.4977
11.0731 | -34.724 | -19.5071
11.0792 | -34.7245 | -19.5163
11.0853 | -34.725 | -19.5252
11.0914 | -34.7255 | -19.5343
11.0975 | -34.726 | -19.5438
11.1037 | -34.7265 | -19.5528
11.1098 | -34.7269 | -19.5621
11.1159 | -34.7274 | -19.5716
11.122 | -34.728 | -19.5811
11.1281 | -34.7284 | -19.5901
11.1342 | -34.7289 | -19.599
11.1403 | -34.7294 | -19.6085
11.1464 | -34.7299 | -19.6176
11.1525 | -34.7304 | -19.6271
11.1586 | -34.7308 | -19.6368
11.1647 | -34.7314 | -19.6459
11.1708 | -34.7318 | -19.655
11.1769 | -34.7323 | -19.6638
11.183 | -34.7328 | -19.6729
11.1891 | -34.7333 | -19.6825
11.1952 | -34.7338 | -19.692
11.2013 | -34.7342 | -19.7012
11.2074 | -34.7347 | -19.7101
11.2135 | -34.7351 | -19.7193
11.2196 | -34.7356 | -19.7285
11.2257 | -34.736 | -19.7375
11.2318 | -34.7364 | -19.747
11.2379 | -34.7369 | -19.756
11.244 | -34.7373 | -19.7651
11.2502 | -34.7378 | -19.7743
11.2563 | -34.7383 | -19.7834
11.2624 | -34.7388 | -19.7927
11.2685 | -34.7392 | -19.8018
11.2746 | -34.7397 | -19.8107
11.2807 | -34.7402 | -19.8203
11.2868 | -34.7406 | -19.8298
11.2929 | -34.7412 | -19.8387
11.299 | -34.7416 | -19.8481
11.3051 | -34.7421 | -19.8572
11.3112 | -34.7425 | -19.8664
11.3173 | -34.7431 | -19.8757
11.3234 | -34.7435 | -19.8851
11.3295 | -34.744 | -19.894
11.3356 | -34.7445 | -19.9031
11.3417 | -34.745 | -19.9125
11.3478 | -34.7454 | -19.9216
11.3539 | -34.7459 | -19.9307
11.36 | -34.7464 | -19.9398
11.3661 | -34.7469 | -19.9489
11.3722 | -34.7473 | -19.9581
11.3783 | -34.7478 | -19.9672
11.3844 | -34.7483 | -19.9764
11.3906 | -34.7487 | -19.9859
11.3967 | -34.7492 | -19.9949
11.4028 | -34.7497 | -20.0042
11.4089 | -34.7502 | -20.0133
11.415 | -34.7506 | -20.0224
11.4211 | -34.7511 | -20.0313
11.4272 | -34.7516 | -20.0406
11.4333 | -34.752 | -20.0498
11.4394 | -34.7525 | -20.0593
11.4455 | -34.753 | -20.0686
11.4516 | -34.7535 | -20.0777
11.4577 | -34.7539 | -20.0871
11.4638 | -34.7545 | -20.0958
11.4699 | -34.7549 | -20.1053
11.476 | -34.7554 | -20.1148
11.4821 | -34.7559 | -20.1237
11.4882 | -34.7564 | -20.1332
11.4943 | -34.7568 | -20.1421
11.5004 | -34.7574 | -20.1513
11.5065 | -34.7579 | -20.1605
11.5126 | -34.7584 | -20.1695
11.5187 | -34.7589 | -20.1786
11.5248 | -34.7593 | -20.1877
11.5309 | -34.7598 | -20.1969
11.5371 | -34.7603 | -20.206
11.5432 | -34.7607 | -20.2154
11.5493 | -34.7612 | -20.2246
11.5554 | -34.7617 | -20.2336
11.5615 | -34.7621 | -20.2429
11.5676 | -34.7627 | -20.2523
11.5737 | -34.7632 | -20.2612
11.5798 | -34.7636 | -20.2702
11.5859 | -34.7641 | -20.2792
11.592 | -34.7646 | -20.2885
11.5981 | -34.7651 | -20.2975
11.6042 | -34.7656 | -20.3068
11.6103 | -34.766 | -20.3159
11.6164 | -34.7665 | -20.3249
11.6225 | -34.767 | -20.334
11.6286 | -34.7675 | -20.3433
11.6347 | -34.768 | -20.3527
11.6408 | -34.7685 | -20.3619
11.6469 | -34.769 | -20.3712
11.653 | -34.7695 | -20.3802
11.6591 | -34.77 | -20.3894
11.6652 | -34.7705 | -20.3987
11.6713 | -34.7709 | -20.4078
11.6775 | -34.7714 | -20.4168
11.6836 | -34.7719 | -20.4261
11.6897 | -34.7724 | -20.435
11.6958 | -34.7729 | -20.4445
11.7019 | -34.7734 | -20.4537
11.708 | -34.7739 | -20.4631
11.7141 | -34.7744 | -20.4722
11.7202 | -34.7749 | -20.4811
11.7263 | -34.7754 | -20.4904
11.7324 | -34.7759 | -20.4996
11.7385 | -34.7763 | -20.5089
11.7446 | -34.7768 | -20.5183
11.7507 | -34.7773 | -20.5274
11.7568 | -34.7778 | -20.5366
11.7629 | -34.7783 | -20.5457
11.769 | -34.7788 | -20.555
11.7751 | -34.7793 | -20.5639
11.7812 | -34.7798 | -20.5729
11.7873 | -34.7803 | -20.5823
11.7934 | -34.7808 | -20.5914
11.7995 | -34.7813 | -20.6003
11.8056 | -34.7817 | -20.6098
11.8117 | -34.7822 | -20.619
11.8178 | -34.7827 | -20.6281
11.824 | -34.7832 | -20.6371
11.8301 | -34.7838 | -20.6459
11.8362 | -34.7842 | -20.6548
11.8423 | -34.7847 | -20.664
11.8484 | -34.7852 | -20.6731
11.8545 | -34.7857 | -20.6824
11.8606 | -34.7862 | -20.6917
11.8667 | -34.7867 | -20.7008
11.8728 | -34.7871 | -20.71
11.8789 | -34.7876 | -20.7196
11.885 | -34.788 | -20.7286
11.8911 | -34.7886 | -20.738
11.8972 | -34.7891 | -20.7467
11.9033 | -34.7895 | -20.7555
11.9094 | -34.79 | -20.7643
11.9155 | -34.7904 | -20.7735
11.9216 | -34.791 | -20.7825
11.9277 | -34.7915 | -20.7914
11.9338 | -34.7919 | -20.8006
11.9399 | -34.7924 | -20.81
11.946 | -34.7929 | -20.8192
11.9521 | -34.7934 | -20.8283
11.9582 | -34.7939 | -20.8372
11.9644 | -34.7944 | -20.8463
11.9705 | -34.7949 | -20.8556
11.9766 | -34.7954 | -20.8647
11.9827 | -34.7959 | -20.8736
11.9888 | -34.7963 | -20.8825
11.9949 | -34.7968 | -20.8916
12.001 | -34.7973 | -20.9008
12.0071 | -34.7978 | -20.9097
12.0132 | -34.7983 | -20.9185
12.0193 | -34.7988 | -20.9279
12.0254 | -34.7993 | -20.9368
12.0315 | -34.7998 | -20.9458
12.0376 | -34.8003 | -20.9547
12.0437 | -34.8008 | -20.9636
12.0498 | -34.8012 | -20.9729
12.0559 | -34.8017 | -20.9821
12.062 | -34.8022 | -20.9913
12.0681 | -34.8027 | -21.0003
12.0742 | -34.8032 | -21.0096
12.0803 | -34.8037 | -21.0186
12.0864 | -34.8042 | -21.0276
12.0925 | -34.8047 | -21.0367
12.0986 | -34.8052 | -21.0455
12.1047 | -34.8058 | -21.0546
12.1109 | -34.8063 | -21.0639
12.117 | -34.8068 | -21.0729
12.1231 | -34.8073 | -21.082
12.1292 | -34.8078 | -21.0913
12.1353 | -34.8083 | -21.1002
12.1414 | -34.8088 | -21.1094
12.1475 | -34.8092 | -21.1185
12.1536 | -34.8097 | -21.1276
12.1597 | -34.8102 | -21.1368
12.1658 | -34.8107 | -21.1458
12.1719 | -34.8112 | -21.1551
12.178 | -34.8118 | -21.1643
12.1841 | -34.8122 | -21.1735
12.1902 | -34.8127 | -21.1826
12.1963 | -34.8132 | -21.1916
12.2024 | -34.8138 | -21.2006
12.2085 | -34.8142 | -21.2097
12.2146 | -34.8147 | -21.2187
12.2207 | -34.8152 | -21.2278
12.2268 | -34.8157 | -21.2368
12.2329 | -34.8162 | -21.2458
12.239 | -34.8166 | -21.2547
12.2451 | -34.8172 | -21.2635
12.2513 | -34.8177 | -21.2726
12.2574 | -34.8182 | -21.2819
12.2635 | -34.8187 | -21.2905
12.2696 | -34.8192 | -21.2994
12.2757 | -34.8196 | -21.3082
12.2818 | -34.8201 | -21.3171
12.2879 | -34.8206 | -21.326
12.294 | -34.8211 | -21.3349
12.3001 | -34.8216 | -21.3441
12.3062 | -34.8221 | -21.3531
12.3123 | -34.8226 | -21.3623
12.3184 | -34.8231 | -21.3712
12.3245 | -34.8236 | -21.3804
12.3306 | -34.8241 | -21.3892
12.3367 | -34.8246 | -21.398
12.3428 | -34.8251 | -21.4068
12.3489 | -34.8256 | -21.4154
12.355 | -34.8261 | -21.4244
12.3611 | -34.8266 | -21.4333
12.3672 | -34.8271 | -21.4422
12.3733 | -34.8276 | -21.4515
12.3794 | -34.828 | -21.46
12.3855 | -34.8285 | -21.469
12.3916 | -34.829 | -21.4782
12.3978 | -34.8295 | -21.4874
12.4039 | -34.8299 | -21.4963
12.41 | -34.8305 | -21.5055
12.4161 | -34.831 | -21.5142
12.4222 | -34.8315 | -21.5231
12.4283 | -34.832 | -21.5321
12.4344 | -34.8325 | -21.541
12.4405 | -34.833 | -21.55
12.4466 | -34.8335 | -21.5589
12.4527 | -34.834 | -21.568
12.4588 | -34.8345 | -21.577
12.4649 | -34.835 | -21.5862
12.471 | -34.8355 | -21.5956
12.4771 | -34.836 | -21.604
12.4832 | -34.8365 | -21.613
12.4893 | -34.837 | -21.6217
12.4954 | -34.8376 | -21.6306
12.5015 | -34.8381 | -21.6393
12.5076 | -34.8385 | -21.6483
12.5137 | -34.839 | -21.6577
12.5198 | -34.8395 | -21.6664
12.5259 | -34.84 | -21.6752
12.532 | -34.8406 | -21.6843
12.5382 | -34.841 | -21.6931
12.5443 | -34.8415 | -21.7022
12.5504 | -34.842 | -21.7112
12.5565 | -34.8426 | -21.7199
12.5626 | -34.8431 | -21.7291
12.5687 | -34.8436 | -21.7382
12.5748 | -34.8441 | -21.747
12.5809 | -34.8446 | -21.7556
12.587 | -34.845 | -21.7649
12.5931 | -34.8455 | -21.7738
12.5992 | -34.846 | -21.783
12.6053 | -34.8465 | -21.7919
12.6114 | -34.8471 | -21.8008
12.6175 | -34.8476 | -21.8099
12.6236 | -34.8481 | -21.8186
12.6297 | -34.8486 | -21.8273
12.6358 | -34.8491 | -21.8361
12.6419 | -34.8497 | -21.8453
12.648 | -34.8501 | -21.8544
12.6541 | -34.8507 | -21.8633
12.6602 | -34.8512 | -21.8719
12.6663 | -34.8516 | -21.8806
12.6724 | -34.8522 | -21.8896
12.6785 | -34.8527 | -21.8987
12.6847 | -34.8532 | -21.9078
12.6908 | -34.8537 | -21.9166
12.6969 | -34.8542 | -21.9255
12.703 | -34.8547 | -21.9344
12.7091 | -34.8552 | -21.9434
12.7152 | -34.8557 | -21.9524
12.7213 | -34.8562 | -21.9614
12.7274 | -34.8567 | -21.9703
12.7335 | -34.8572 | -21.9787
12.7396 | -34.8577 | -21.9873
12.7457 | -34.8582 | -21.9962
12.7518 | -34.8587 | -22.0051
12.7579 | -34.8591 | -22.0142
12.764 | -34.8596 | -22.0231
12.7701 | -34.8601 | -22.0317
12.7762 | -34.8606 | -22.0407
12.7823 | -34.861 | -22.0496
12.7884 | -34.8615 | -22.0583
12.7945 | -34.862 | -22.0667
12.8006 | -34.8625 | -22.0748
12.8067 | -34.863 | -22.084
12.8128 | -34.8635 | -22.0926
12.8189 | -34.8639 | -22.1012
12.8251 | -34.8644 | -22.1099
12.8312 | -34.8649 | -22.1187
12.8373 | -34.8654 | -22.1275
12.8434 | -34.8659 | -22.1362
12.8495 | -34.8663 | -22.145
12.8556 | -34.8668 | -22.1538
12.8617 | -34.8673 | -22.1624
12.8678 | -34.8678 | -22.1712
12.8739 | -34.8683 | -22.1796
12.88 | -34.8688 | -22.1885
12.8861 | -34.8692 | -22.1975
12.8922 | -34.8697 | -22.2062
12.8983 | -34.8703 | -22.2146
12.9044 | -34.8708 | -22.2234
12.9105 | -34.8713 | -22.2321
12.9166 | -34.8718 | -22.2414
12.9227 | -34.8723 | -22.2507
12.9288 | -34.8729 | -22.2597
12.9349 | -34.8734 | -22.2689
12.941 | -34.8739 | -22.2778
12.9471 | -34.8744 | -22.2866
12.9532 | -34.875 | -22.2955
12.9593 | -34.8755 | -22.3044
12.9654 | -34.8759 | -22.3133
12.9716 | -34.8764 | -22.322
12.9777 | -34.8769 | -22.3301
12.9838 | -34.8773 | -22.3388
12.9899 | -34.8778 | -22.3478
12.996 | -34.8783 | -22.3568
13.0021 | -34.8787 | -22.3655
13.0082 | -34.8792 | -22.3742
13.0143 | -34.8797 | -22.3826
13.0204 | -34.8801 | -22.3911
13.0265 | -34.8807 | -22.3997
13.0326 | -34.8811 | -22.4087
13.0387 | -34.8816 | -22.4175
13.0448 | -34.8821 | -22.4262
13.0509 | -34.8826 | -22.4349
13.057 | -34.8832 | -22.4434
13.0631 | -34.8837 | -22.452
13.0692 | -34.8842 | -22.4606
13.0753 | -34.8848 | -22.4693
13.0814 | -34.8852 | -22.4776
13.0875 | -34.8857 | -22.4859
13.0936 | -34.8862 | -22.4945
13.0997 | -34.8867 | -22.5028
13.1058 | -34.8872 | -22.5109
13.112 | -34.8876 | -22.5196
13.1181 | -34.8881 | -22.5282
13.1242 | -34.8886 | -22.537
13.1303 | -34.8891 | -22.5459
13.1364 | -34.8896 | -22.5547
13.1425 | -34.8901 | -22.5634
13.1486 | -34.8906 | -22.572
13.1547 | -34.891 | -22.5804
13.1608 | -34.8915 | -22.5889
13.1669 | -34.892 | -22.5975
13.173 | -34.8925 | -22.6061
13.1791 | -34.893 | -22.6148
13.1852 | -34.8935 | -22.6238
13.1913 | -34.894 | -22.6324
13.1974 | -34.8944 | -22.6411
13.2035 | -34.8949 | -22.6498
13.2096 | -34.8954 | -22.6584
13.2157 | -34.8959 | -22.6671
13.2218 | -34.8965 | -22.676
13.2279 | -34.897 | -22.6846
13.234 | -34.8975 | -22.6932
13.2401 | -34.8979 | -22.7017
13.2462 | -34.8984 | -22.7102
13.2524 | -34.8989 | -22.7188
13.2585 | -34.8994 | -22.7273
13.2646 | -34.8999 | -22.7363
13.2707 | -34.9004 | -22.745
13.2768 | -34.9009 | -22.7538
13.2829 | -34.9013 | -22.7625
13.289 | -34.9018 | -22.7711
13.2951 | -34.9023 | -22.7794
13.3012 | -34.9028 | -22.7881
13.3073 | -34.9032 | -22.7969
13.3134 | -34.9037 | -22.8054
13.3195 | -34.9042 | -22.814
13.3256 | -34.9047 | -22.823
13.3317 | -34.9052 | -22.8315
13.3378 | -34.9057 | -22.8402
13.3439 | -34.9062 | -22.8487
13.35 | -34.9067 | -22.8574
13.3561 | -34.9072 | -22.866
13.3622 | -34.9076 | -22.8745
13.3683 | -34.9082 | -22.883
13.3744 | -34.9086 | -22.8916
13.3805 | -34.9091 | -22.9001
13.3866 | -34.9095 | -22.9087
13.3927 | -34.91 | -22.9173
13.3989 | -34.9105 | -22.9256
13.405 | -34.9111 | -22.9341
13.4111 | -34.9116 | -22.9427
13.4172 | -34.9121 | -22.951
13.4233 | -34.9125 | -22.9598
13.4294 | -34.913 | -22.9681
13.4355 | -34.9136 | -22.977
13.4416 | -34.914 | -22.9857
13.4477 | -34.9144 | -22.9944
13.4538 | -34.9149 | -23.0028
13.4599 | -34.9154 | -23.0115
13.466 | -34.9159 | -23.0205
13.4721 | -34.9165 | -23.0291
13.4782 | -34.9169 | -23.0375
13.4843 | -34.9175 | -23.0461
13.4904 | -34.918 | -23.0541
13.4965 | -34.9184 | -23.0629
13.5026 | -34.9189 | -23.0714
13.5087 | -34.9195 | -23.0792
13.5148 | -34.9199 | -23.0877
13.5209 | -34.9204 | -23.096
13.527 | -34.9209 | -23.105
13.5331 | -34.9213 | -23.1134
13.5393 | -34.9218 | -23.122
13.5454 | -34.9222 | -23.1307
13.5515 | -34.9228 | -23.1391
13.5576 | -34.9233 | -23.1478
13.5637 | -34.9238 | -23.1563
13.5698 | -34.9243 | -23.1651
13.5759 | -34.9247 | -23.1734
13.582 | -34.9251 | -23.1822
13.5881 | -34.9256 | -23.1906
13.5942 | -34.9261 | -23.1992
13.6003 | -34.9266 | -23.2077
13.6064 | -34.9271 | -23.2165
13.6125 | -34.9276 | -23.2254
13.6186 | -34.9282 | -23.234
13.6247 | -34.9286 | -23.2424
13.6308 | -34.9291 | -23.2506
13.6369 | -34.9296 | -23.2592
13.643 | -34.9301 | -23.2674
13.6491 | -34.9306 | -23.2759
13.6552 | -34.9311 | -23.2842
13.6613 | -34.9315 | -23.2931
13.6674 | -34.9321 | -23.3019
13.6735 | -34.9325 | -23.3107
13.6796 | -34.933 | -23.319
13.6858 | -34.9334 | -23.3277
13.6919 | -34.9339 | -23.3363
13.698 | -34.9344 | -23.3449
13.7041 | -34.9349 | -23.3536
13.7102 | -34.9354 | -23.3622
13.7163 | -34.9359 | -23.3708
13.7224 | -34.9365 | -23.3792
13.7285 | -34.9369 | -23.3878
13.7346 | -34.9374 | -23.3964
13.7407 | -34.9379 | -23.405
13.7468 | -34.9384 | -23.4136
13.7529 | -34.9388 | -23.422
13.759 | -34.9393 | -23.4306
13.7651 | -34.9398 | -23.4392
13.7712 | -34.9403 | -23.4479
13.7773 | -34.9408 | -23.4566
13.7834 | -34.9413 | -23.4651
13.7895 | -34.9419 | -23.4735
13.7956 | -34.9424 | -23.4823
13.8017 | -34.9429 | -23.4908
13.8078 | -34.9434 | -23.4994
13.8139 | -34.9439 | -23.5077
13.82 | -34.9444 | -23.5161
13.8262 | -34.9449 | -23.5247
13.8323 | -34.9454 | -23.5333
13.8384 | -34.9459 | -23.5419
13.8445 | -34.9464 | -23.5505
13.8506 | -34.9469 | -23.5591
13.8567 | -34.9474 | -23.5677
13.8628 | -34.9479 | -23.5764
13.8689 | -34.9483 | -23.5851
13.875 | -34.9488 | -23.5936
13.8811 | -34.9494 | -23.6018
13.8872 | -34.9499 | -23.6102
13.8933 | -34.9504 | -23.6187
13.8994 | -34.9509 | -23.627
13.9055 | -34.9514 | -23.6352
13.9116 | -34.9519 | -23.6434
13.9177 | -34.9524 | -23.6514
13.9238 | -34.953 | -23.6603
13.9299 | -34.9535 | -23.6687
13.936 | -34.954 | -23.6769
13.9421 | -34.9544 | -23.6854
13.9482 | -34.955 | -23.6939
13.9543 | -34.9555 | -23.7021
13.9604 | -34.9559 | -23.7105
13.9665 | -34.9564 | -23.7193
13.9727 | -34.9569 | -23.7273
13.9788 | -34.9573 | -23.7362
13.9849 | -34.9578 | -23.7447
13.991 | -34.9582 | -23.7534
13.9971 | -34.9587 | -23.7623
14.0032 | -34.9591 | -23.771
14.0093 | -34.9596 | -23.7792
14.0154 | -34.9601 | -23.7879
14.0215 | -34.9605 | -23.7961
14.0276 | -34.961 | -23.8044
14.0337 | -34.9616 | -23.8131
14.0398 | -34.9621 | -23.8215
14.0459 | -34.9625 | -23.8303
14.052 | -34.9629 | -23.8388
14.0581 | -34.9634 | -23.847
14.0642 | -34.964 | -23.8558
14.0703 | -34.9645 | -23.8641
14.0764 | -34.9649 | -23.8731
14.0825 | -34.9654 | -23.8816
14.0886 | -34.9659 | -23.8896
14.0947 | -34.9663 | -23.8978
14.1008 | -34.9668 | -23.9067
14.1069 | -34.9673 | -23.915
14.1131 | -34.9678 | -23.9238
14.1192 | -34.9682 | -23.932
14.1253 | -34.9687 | -23.9402
14.1314 | -34.9692 | -23.9486
14.1375 | -34.9697 | -23.9573
14.1436 | -34.9703 | -23.9658
14.1497 | -34.9708 | -23.974
14.1558 | -34.9713 | -23.9826
14.1619 | -34.9718 | -23.9911
14.168 | -34.9723 | -23.9994
14.1741 | -34.9728 | -24.0077
14.1802 | -34.9733 | -24.0161
14.1863 | -34.9738 | -24.0246
14.1924 | -34.9743 | -24.033
14.1985 | -34.9747 | -24.0413
14.2046 | -34.9753 | -24.0494
14.2107 | -34.9758 | -24.0573
14.2168 | -34.9763 | -24.0661
14.2229 | -34.9768 | -24.0746
14.229 | -34.9773 | -24.0831
14.2351 | -34.9778 | -24.0914
14.2412 | -34.9783 | -24.0995
14.2473 | -34.9788 | -24.108
14.2534 | -34.9793 | -24.1165
14.2596 | -34.9799 | -24.1249
14.2657 | -34.9804 | -24.1336
14.2718 | -34.9808 | -24.1421
14.2779 | -34.9813 | -24.1501
14.284 | -34.9818 | -24.1582
14.2901 | -34.9824 | -24.1665
14.2962 | -34.9829 | -24.1746
14.3023 | -34.9834 | -24.1832
14.3084 | -34.9839 | -24.1917
14.3145 | -34.9844 | -24.2002
14.3206 | -34.9849 | -24.2085
14.3267 | -34.9854 | -24.2168
14.3328 | -34.9859 | -24.2251
14.3389 | -34.9863 | -24.2334
14.345 | -34.9868 | -24.242
14.3511 | -34.9873 | -24.2501
14.3572 | -34.9878 | -24.2587
14.3633 | -34.9883 | -24.2669
14.3694 | -34.9887 | -24.2752
14.3755 | -34.9892 | -24.2832
14.3816 | -34.9897 | -24.2914
14.3877 | -34.9902 | -24.2995
14.3938 | -34.9907 | -24.3079
14.4 | -34.9912 | -24.3161
14.4061 | -34.9916 | -24.3245
14.4122 | -34.9921 | -24.3331
14.4183 | -34.9926 | -24.3416
14.4244 | -34.9931 | -24.3497
14.4305 | -34.9936 | -24.3584
14.4366 | -34.9941 | -24.3668
14.4427 | -34.9946 | -24.3753
14.4488 | -34.9951 | -24.3833
14.4549 | -34.9956 | -24.3916
14.461 | -34.9961 | -24.4001
14.4671 | -34.9966 | -24.4086
14.4732 | -34.9971 | -24.4165
14.4793 | -34.9976 | -24.4255
14.4854 | -34.9982 | -24.4334
14.4915 | -34.9987 | -24.4415
14.4976 | -34.9992 | -24.4499
14.5037 | -34.9997 | -24.4584
14.5098 | -35.0002 | -24.4665
14.5159 | -35.0007 | -24.475
14.522 | -35.0012 | -24.4835
14.5281 | -35.0018 | -24.4915
14.5342 | -35.0022 | -24.5
14.5403 | -35.0028 | -24.5084
14.5465 | -35.0033 | -24.517
14.5526 | -35.0038 | -24.5255
14.5587 | -35.0042 | -24.5338
14.5648 | -35.0047 | -24.5418
14.5709 | -35.0052 | -24.5504
14.577 | -35.0057 | -24.5588
14.5831 | -35.0063 | -24.5673
14.5892 | -35.0068 | -24.5752
14.5953 | -35.0073 | -24.5834
14.6014 | -35.0078 | -24.5917
14.6075 | -35.0083 | -24.5999
14.6136 | -35.0088 | -24.608
14.6197 | -35.0093 | -24.6165
14.6258 | -35.0098 | -24.6249
14.6319 | -35.0103 | -24.6328
14.638 | -35.0109 | -24.6411
14.6441 | -35.0114 | -24.6488
14.6502 | -35.0119 | -24.6567
14.6563 | -35.0124 | -24.665
14.6624 | -35.0129 | -24.6731
14.6685 | -35.0134 | -24.681
14.6746 | -35.0139 | -24.6893
14.6807 | -35.0144 | -24.6976
14.6869 | -35.0149 | -24.7058
14.693 | -35.0154 | -24.714
14.6991 | -35.0158 | -24.7221
14.7052 | -35.0164 | -24.7303
14.7113 | -35.0169 | -24.7387
14.7174 | -35.0174 | -24.7468
14.7235 | -35.0179 | -24.755
14.7296 | -35.0184 | -24.7632
14.7357 | -35.0189 | -24.7711
14.7418 | -35.0194 | -24.7791
14.7479 | -35.0199 | -24.7874
14.754 | -35.0205 | -24.7956
14.7601 | -35.021 | -24.8039
14.7662 | -35.0215 | -24.8121
14.7723 | -35.022 | -24.8202
14.7784 | -35.0225 | -24.8286
14.7845 | -35.023 | -24.8368
14.7906 | -35.0235 | -24.845
14.7967 | -35.024 | -24.8531
14.8028 | -35.0245 | -24.8611
14.8089 | -35.025 | -24.869
14.815 | -35.0254 | -24.8769
14.8211 | -35.0259 | -24.8849
14.8272 | -35.0264 | -24.8932
14.8334 | -35.0269 | -24.9013
14.8395 | -35.0274 | -24.9094
14.8456 | -35.028 | -24.9174
14.8517 | -35.0285 | -24.9256
14.8578 | -35.029 | -24.9336
14.8639 | -35.0295 | -24.942
14.87 | -35.03 | -24.9502
14.8761 | -35.0305 | -24.9579
14.8822 | -35.031 | -24.966
14.8883 | -35.0315 | -24.974
14.8944 | -35.032 | -24.9823
14.9005 | -35.0326 | -24.9902
14.9066 | -35.033 | -24.9983
14.9127 | -35.0336 | -25.0062
14.9188 | -35.0341 | -25.0146
14.9249 | -35.0346 | -25.0228
14.931 | -35.0351 | -25.0307
14.9371 | -35.0356 | -25.0388
14.9432 | -35.0361 | -25.047
14.9493 | -35.0366 | -25.0551
14.9554 | -35.0371 | -25.0631
14.9615 | -35.0376 | -25.0714
14.9676 | -35.0382 | -25.0798
14.9738 | -35.0387 | -25.0878
14.9799 | -35.0392 | -25.0962
14.986 | -35.0397 | -25.1039
14.9921 | -35.0403 | -25.1119
14.9982 | -35.0407 | -25.1203
15.0043 | -35.0412 | -25.1283
15.0104 | -35.0417 | -25.1365
15.0165 | -35.0423 | -25.1447
15.0226 | -35.0428 | -25.1524
15.0287 | -35.0433 | -25.161
15.0348 | -35.0438 | -25.1691
15.0409 | -35.0443 | -25.1776
15.047 | -35.0448 | -25.1859
15.0531 | -35.0454 | -25.1941
15.0592 | -35.046 | -25.2023
15.0653 | -35.0465 | -25.2104
15.0714 | -35.047 | -25.2183
15.0775 | -35.0475 | -25.2261
15.0836 | -35.048 | -25.2342
15.0897 | -35.0486 | -25.2424
15.0958 | -35.049 | -25.2503
15.1019 | -35.0495 | -25.2585
15.108 | -35.05 | -25.2667
15.1141 | -35.0505 | -25.2748
15.1203 | -35.0511 | -25.2827
15.1264 | -35.0517 | -25.2909
15.1325 | -35.0522 | -25.2989
15.1386 | -35.0527 | -25.3071
15.1447 | -35.0532 | -25.3152
15.1508 | -35.0538 | -25.3235
15.1569 | -35.0543 | -25.3318
15.163 | -35.0549 | -25.3394
15.1691 | -35.0554 | -25.3474
15.1752 | -35.0559 | -25.3555
15.1813 | -35.0565 | -25.3636
15.1874 | -35.057 | -25.3718
15.1935 | -35.0576 | -25.3796
15.1996 | -35.0581 | -25.3879
15.2057 | -35.0586 | -25.3958
15.2118 | -35.0591 | -25.4038
15.2179 | -35.0597 | -25.4116
15.224 | -35.0601 | -25.4197
15.2301 | -35.0607 | -25.4273
15.2362 | -35.0612 | -25.4355
15.2423 | -35.0618 | -25.4438
15.2484 | -35.0623 | -25.4517
15.2545 | -35.0628 | -25.4597
15.2607 | -35.0633 | -25.468
15.2668 | -35.0639 | -25.4757
15.2729 | -35.0644 | -25.4837
15.279 | -35.0649 | -25.4917
15.2851 | -35.0654 | -25.4998
15.2912 | -35.066 | -25.5076
15.2973 | -35.0664 | -25.5155
15.3034 | -35.067 | -25.5233
15.3095 | -35.0675 | -25.531
15.3156 | -35.068 | -25.5392
15.3217 | -35.0685 | -25.5468
15.3278 | -35.0691 | -25.5548
15.3339 | -35.0696 | -25.5626
15.34 | -35.0702 | -25.5704
15.3461 | -35.0706 | -25.5786
15.3522 | -35.0711 | -25.5865
15.3583 | -35.0717 | -25.5944
15.3644 | -35.0722 | -25.6025
15.3705 | -35.0727 | -25.6103
15.3766 | -35.0733 | -25.6186
15.3827 | -35.0738 | -25.6263
15.3888 | -35.0743 | -25.6341
15.3949 | -35.0748 | -25.642
15.401 | -35.0754 | -25.65
15.4072 | -35.0759 | -25.6578
15.4133 | -35.0764 | -25.6657
15.4194 | -35.077 | -25.6738
15.4255 | -35.0775 | -25.6817
15.4316 | -35.0781 | -25.6896
15.4377 | -35.0786 | -25.6975
15.4438 | -35.0791 | -25.7049
15.4499 | -35.0796 | -25.7129
15.456 | -35.0801 | -25.7209
15.4621 | -35.0806 | -25.7288
15.4682 | -35.0812 | -25.7369
15.4743 | -35.0818 | -25.7446
15.4804 | -35.0823 | -25.7527
15.4865 | -35.0828 | -25.7607
15.4926 | -35.0833 | -25.7688
15.4987 | -35.0839 | -25.7769
15.5048 | -35.0844 | -25.7847
15.5109 | -35.085 | -25.7923
15.517 | -35.0855 | -25.8003
15.5231 | -35.086 | -25.8081
15.5292 | -35.0866 | -25.8159
15.5353 | -35.0871 | -25.8238
15.5414 | -35.0877 | -25.832
15.5476 | -35.0882 | -25.8398
15.5537 | -35.0887 | -25.8476
15.5598 | -35.0893 | -25.8553
15.5659 | -35.0898 | -25.8634
15.572 | -35.0904 | -25.871
15.5781 | -35.0909 | -25.879
15.5842 | -35.0914 | -25.8866
15.5903 | -35.092 | -25.8942
15.5964 | -35.0925 | -25.9019
15.6025 | -35.0931 | -25.9096
15.6086 | -35.0936 | -25.9174
15.6147 | -35.0941 | -25.9253
15.6208 | -35.0946 | -25.9329
15.6269 | -35.0952 | -25.9407
15.633 | -35.0957 | -25.9485
15.6391 | -35.0962 | -25.9566
15.6452 | -35.0968 | -25.9642
15.6513 | -35.0973 | -25.9724
15.6574 | -35.0979 | -25.9799
15.6635 | -35.0984 | -25.9879
15.6696 | -35.0989 | -25.9957
15.6757 | -35.0995 | -26.0039
15.6818 | -35.1 | -26.0121
15.688 | -35.1006 | -26.0198
15.6941 | -35.1011 | -26.0277
15.7002 | -35.1017 | -26.0356
15.7063 | -35.1022 | -26.0436
15.7124 | -35.1027 | -26.0516
15.7185 | -35.1033 | -26.0594
15.7246 | -35.1038 | -26.0671
15.7307 | -35.1044 | -26.075
15.7368 | -35.1049 | -26.083
15.7429 | -35.1054 | -26.0909
15.749 | -35.106 | -26.0986
15.7551 | -35.1065 | -26.1064
15.7612 | -35.107 | -26.1135
15.7673 | -35.1075 | -26.1212
15.7734 | -35.108 | -26.1288
15.7795 | -35.1086 | -26.137
15.7856 | -35.1091 | -26.1446
15.7917 | -35.1096 | -26.1522
15.7978 | -35.1102 | -26.16
15.8039 | -35.1107 | -26.1678
15.81 | -35.1112 | -26.1753
15.8161 | -35.1117 | -26.1833
15.8222 | -35.1123 | -26.1911
15.8283 | -35.1129 | -26.1988
15.8345 | -35.1134 | -26.2066
15.8406 | -35.114 | -26.2145
15.8467 | -35.1145 | -26.2222
15.8528 | -35.1151 | -26.2298
15.8589 | -35.1156 | -26.2375
15.865 | -35.1161 | -26.2451
15.8711 | -35.1166 | -26.2532
15.8772 | -35.1172 | -26.261
15.8833 | -35.1178 | -26.2692
15.8894 | -35.1183 | -26.2768
15.8955 | -35.1188 | -26.2848
15.9016 | -35.1194 | -26.2924
15.9077 | -35.1199 | -26.3002
15.9138 | -35.1205 | -26.3079
15.9199 | -35.121 | -26.3159
15.926 | -35.1216 | -26.3241
15.9321 | -35.1222 | -26.3318
15.9382 | -35.1227 | -26.3395
15.9443 | -35.1233 | -26.3476
15.9504 | -35.1238 | -26.3554
15.9565 | -35.1244 | -26.3632
15.9626 | -35.125 | -26.371
15.9687 | -35.1255 | -26.3785
15.9749 | -35.1261 | -26.3864
15.981 | -35.1267 | -26.3942
15.9871 | -35.1272 | -26.402
15.9932 | -35.1277 | -26.4099
15.9993 | -35.1283 | -26.4176
16.0054 | -35.1288 | -26.4251
16.0115 | -35.1294 | -26.4326
16.0176 | -35.13 | -26.4406
16.0237 | -35.1305 | -26.4487
16.0298 | -35.1311 | -26.456
16.0359 | -35.1316 | -26.4637
16.042 | -35.1322 | -26.4715
16.0481 | -35.1328 | -26.4792
16.0542 | -35.1333 | -26.4872
16.0603 | -35.1339 | -26.4949
16.0664 | -35.1344 | -26.5027
16.0725 | -35.135 | -26.5103
16.0786 | -35.1356 | -26.5184
16.0847 | -35.1361 | -26.5259
16.0908 | -35.1367 | -26.5335
16.0969 | -35.1372 | -26.541
16.103 | -35.1377 | -26.5489
16.1091 | -35.1383 | -26.5565
16.1152 | -35.1388 | -26.5643
16.1214 | -35.1394 | -26.572
16.1275 | -35.14 | -26.5796
16.1336 | -35.1405 | -26.5873
16.1397 | -35.1411 | -26.5951
16.1458 | -35.1416 | -26.6028
16.1519 | -35.1422 | -26.6107
16.158 | -35.1427 | -26.6185
16.1641 | -35.1433 | -26.6262
16.1702 | -35.1438 | -26.634
16.1763 | -35.1444 | -26.6418
16.1824 | -35.145 | -26.6493
16.1885 | -35.1455 | -26.6571
16.1946 | -35.1461 | -26.6649
16.2007 | -35.1466 | -26.6731
16.2068 | -35.1472 | -26.6806
16.2129 | -35.1477 | -26.6884
16.219 | -35.1482 | -26.6961
16.2251 | -35.1487 | -26.7041
16.2312 | -35.1492 | -26.7121
16.2373 | -35.1498 | -26.7198
16.2434 | -35.1503 | -26.7276
16.2495 | -35.1509 | -26.7351
16.2556 | -35.1515 | -26.7428
16.2618 | -35.152 | -26.7504
16.2679 | -35.1526 | -26.7581
16.274 | -35.1531 | -26.766
16.2801 | -35.1536 | -26.7736
16.2862 | -35.1542 | -26.7811
16.2923 | -35.1547 | -26.7891
16.2984 | -35.1553 | -26.797
16.3045 | -35.1559 | -26.8048
16.3106 | -35.1564 | -26.8122
16.3167 | -35.157 | -26.8201
16.3228 | -35.1575 | -26.8278
16.3289 | -35.1581 | -26.8354
16.335 | -35.1586 | -26.8434
16.3411 | -35.1591 | -26.8511
16.3472 | -35.1597 | -26.8589
16.3533 | -35.1602 | -26.8667
16.3594 | -35.1608 | -26.8741
16.3655 | -35.1613 | -26.8819
16.3716 | -35.1618 | -26.8896
16.3777 | -35.1624 | -26.8973
16.3838 | -35.1629 | -26.905
16.3899 | -35.1635 | -26.9127
16.396 | -35.164 | -26.9205
16.4021 | -35.1646 | -26.9282
16.4083 | -35.1652 | -26.9359
16.4144 | -35.1658 | -26.9434
16.4205 | -35.1663 | -26.9511
16.4266 | -35.1669 | -26.9588
16.4327 | -35.1675 | -26.9665
16.4388 | -35.168 | -26.9742
16.4449 | -35.1686 | -26.9818
16.451 | -35.1691 | -26.9896
16.4571 | -35.1697 | -26.9974
16.4632 | -35.1702 | -27.0049
16.4693 | -35.1708 | -27.0127
16.4754 | -35.1713 | -27.0204
16.4815 | -35.1719 | -27.0282
16.4876 | -35.1724 | -27.0359
16.4937 | -35.173 | -27.0436
16.4998 | -35.1736 | -27.0512
16.5059 | -35.1742 | -27.0588
16.512 | -35.1747 | -27.0667
16.5181 | -35.1753 | -27.0746
16.5242 | -35.1758 | -27.0825
16.5303 | -35.1764 | -27.0901
16.5364 | -35.1769 | -27.0976
16.5425 | -35.1774 | -27.1053
16.5487 | -35.178 | -27.1131
16.5548 | -35.1785 | -27.1208
16.5609 | -35.1791 | -27.1288
16.567 | -35.1796 | -27.1365
16.5731 | -35.1802 | -27.1441
16.5792 | -35.1807 | -27.1519
16.5853 | -35.1813 | -27.1596
16.5914 | -35.1819 | -27.1674
16.5975 | -35.1824 | -27.1753
16.6036 | -35.183 | -27.183
16.6097 | -35.1835 | -27.1906
16.6158 | -35.1841 | -27.1985
16.6219 | -35.1846 | -27.2061
16.628 | -35.1851 | -27.2137
16.6341 | -35.1857 | -27.2212
16.6402 | -35.1863 | -27.2284
16.6463 | -35.1868 | -27.236
16.6524 | -35.1874 | -27.2438
16.6585 | -35.1879 | -27.2516
16.6646 | -35.1885 | -27.2594
16.6707 | -35.189 | -27.2667
16.6768 | -35.1896 | -27.2742
16.6829 | -35.1901 | -27.2819
16.689 | -35.1907 | -27.2897
16.6952 | -35.1913 | -27.2976
16.7013 | -35.1918 | -27.3052
16.7074 | -35.1923 | -27.3128
16.7135 | -35.1929 | -27.3206
16.7196 | -35.1934 | -27.3281
16.7257 | -35.194 | -27.3358
16.7318 | -35.1945 | -27.3435
16.7379 | -35.1951 | -27.3514
16.744 | -35.1957 | -27.3589
16.7501 | -35.1962 | -27.3666
16.7562 | -35.1968 | -27.3743
16.7623 | -35.1973 | -27.382
16.7684 | -35.1979 | -27.3897
16.7745 | -35.1985 | -27.3971
16.7806 | -35.199 | -27.4049
16.7867 | -35.1996 | -27.4125
16.7928 | -35.2001 | -27.4203
16.7989 | -35.2006 | -27.4281
16.805 | -35.2012 | -27.4358
16.8111 | -35.2018 | -27.4436
16.8172 | -35.2023 | -27.4516
16.8233 | -35.2029 | -27.4597
16.8294 | -35.2034 | -27.4674
16.8356 | -35.2039 | -27.4751
16.8417 | -35.2045 | -27.4827
16.8478 | -35.205 | -27.4903
16.8539 | -35.2056 | -27.498
16.86 | -35.2062 | -27.5055
16.8661 | -35.2067 | -27.5132
16.8722 | -35.2073 | -27.521
16.8783 | -35.2079 | -27.5283
16.8844 | -35.2084 | -27.536
16.8905 | -35.209 | -27.5437
16.8966 | -35.2095 | -27.5515
16.9027 | -35.2101 | -27.559
16.9088 | -35.2106 | -27.5672
16.9149 | -35.2112 | -27.5749
16.921 | -35.2118 | -27.5826
16.9271 | -35.2124 | -27.5904
16.9332 | -35.2129 | -27.5981
16.9393 | -35.2135 | -27.606
16.9454 | -35.214 | -27.6135
16.9515 | -35.2145 | -27.6213
16.9576 | -35.2151 | -27.629
16.9637 | -35.2157 | -27.6369
16.9698 | -35.2163 | -27.6445
16.9759 | -35.2168 | -27.6522
16.9821 | -35.2174 | -27.6598
16.9882 | -35.2179 | -27.6677
16.9943 | -35.2185 | -27.6755
17.0004 | -35.219 | -27.6831
17.0065 | -35.2196 | -27.6908
17.0126 | -35.2201 | -27.6985
17.0187 | -35.2207 | -27.7062
17.0248 | -35.2213 | -27.7139
17.0309 | -35.2218 | -27.7218
17.037 | -35.2224 | -27.7296
17.0431 | -35.223 | -27.7372
17.0492 | -35.2235 | -27.745
17.0553 | -35.224 | -27.7528
17.0614 | -35.2246 | -27.7606
17.0675 | -35.2251 | -27.7686
17.0736 | -35.2257 | -27.7764
17.0797 | -35.2262 | -27.7839
17.0858 | -35.2268 | -27.7918
17.0919 | -35.2273 | -27.7996
17.098 | -35.2279 | -27.8072
17.1041 | -35.2284 | -27.8148
17.1102 | -35.229 | -27.8227
17.1163 | -35.2296 | -27.8306
17.1225 | -35.2301 | -27.8382
17.1286 | -35.2307 | -27.8456
17.1347 | -35.2312 | -27.8531
17.1408 | -35.2318 | -27.8605
17.1469 | -35.2324 | -27.868
17.153 | -35.233 | -27.8756
17.1591 | -35.2335 | -27.8831
17.1652 | -35.234 | -27.8907
17.1713 | -35.2346 | -27.8983
17.1774 | -35.2351 | -27.9062
17.1835 | -35.2357 | -27.9141
17.1896 | -35.2362 | -27.9218
17.1957 | -35.2367 | -27.9296
17.2018 | -35.2373 | -27.9375
17.2079 | -35.2378 | -27.945
17.214 | -35.2384 | -27.9527
17.2201 | -35.2389 | -27.9604
17.2262 | -35.2395 | -27.9681
17.2323 | -35.24 | -27.9758
17.2384 | -35.2406 | -27.9834
17.2445 | -35.2411 | -27.991
17.2506 | -35.2417 | -27.9986
17.2567 | -35.2422 | -28.0061
17.2628 | -35.2428 | -28.0141
17.269 | -35.2433 | -28.0218
17.2751 | -35.2439 | -28.0296
17.2812 | -35.2445 | -28.0371
17.2873 | -35.245 | -28.0449
17.2934 | -35.2455 | -28.0525
17.2995 | -35.246 | -28.06
17.3056 | -35.2466 | -28.0676
17.3117 | -35.2471 | -28.0752
17.3178 | -35.2477 | -28.0829
17.3239 | -35.2482 | -28.0906
17.33 | -35.2488 | -28.0982
17.3361 | -35.2493 | -28.1059
17.3422 | -35.2499 | -28.1136
17.3483 | -35.2505 | -28.1214
17.3544 | -35.251 | -28.1289
17.3605 | -35.2516 | -28.1366
17.3666 | -35.2522 | -28.1443
17.3727 | -35.2527 | -28.1518
17.3788 | -35.2532 | -28.1593
17.3849 | -35.2538 | -28.167
17.391 | -35.2544 | -28.1746
17.3971 | -35.2549 | -28.1824
17.4032 | -35.2555 | -28.1903
17.4094 | -35.256 | -28.198
17.4155 | -35.2565 | -28.2056
17.4216 | -35.2571 | -28.2133
17.4277 | -35.2576 | -28.2213
17.4338 | -35.2582 | -28.229
17.4399 | -35.2588 | -28.2367
17.446 | -35.2593 | -28.2441
17.4521 | -35.2599 | -28.2517
17.4582 | -35.2604 | -28.2591
17.4643 | -35.261 | -28.2671
17.4704 | -35.2615 | -28.2747
17.4765 | -35.262 | -28.2825
17.4826 | -35.2626 | -28.2901
17.4887 | -35.2631 | -28.298
17.4948 | -35.2637 | -28.3056
17.5009 | -35.2642 | -28.3131
17.507 | -35.2648 | -28.3208
17.5131 | -35.2653 | -28.3285
17.5192 | -35.2659 | -28.3362
17.5253 | -35.2664 | -28.344
17.5314 | -35.2669 | -28.3518
17.5375 | -35.2675 | -28.3594
17.5436 | -35.268 | -28.3672
17.5497 | -35.2686 | -28.375
17.5559 | -35.2691 | -28.3825
17.562 | -35.2697 | -28.3899
17.5681 | -35.2702 | -28.3975
17.5742 | -35.2708 | -28.4054
17.5803 | -35.2713 | -28.4128
17.5864 | -35.2719 | -28.4201
17.5925 | -35.2725 | -28.4278
17.5986 | -35.2731 | -28.4352
17.6047 | -35.2737 | -28.4431
17.6108 | -35.2742 | -28.4505
17.6169 | -35.2748 | -28.4581
17.623 | -35.2753 | -28.4657
17.6291 | -35.2759 | -28.4732
17.6352 | -35.2764 | -28.4807
17.6413 | -35.277 | -28.4885
17.6474 | -35.2775 | -28.4963
17.6535 | -35.2781 | -28.504
17.6596 | -35.2786 | -28.5116
17.6657 | -35.2791 | -28.5191
17.6718 | -35.2797 | -28.527
17.6779 | -35.2802 | -28.5346
17.684 | -35.2808 | -28.5422
17.6901 | -35.2813 | -28.5498
17.6963 | -35.2819 | -28.5573
17.7024 | -35.2824 | -28.5649
17.7085 | -35.2829 | -28.5724
17.7146 | -35.2835 | -28.58
17.7207 | -35.284 | -28.5875
17.7268 | -35.2846 | -28.5949
17.7329 | -35.2851 | -28.6027
17.739 | -35.2857 | -28.6104
17.7451 | -35.2863 | -28.6179
17.7512 | -35.2869 | -28.6257
17.7573 | -35.2874 | -28.6334
17.7634 | -35.2879 | -28.641
17.7695 | -35.2885 | -28.6488
17.7756 | -35.289 | -28.6565
17.7817 | -35.2896 | -28.6638
17.7878 | -35.2901 | -28.6715
17.7939 | -35.2907 | -28.6791
17.8 | -35.2912 | -28.6866
17.8061 | -35.2918 | -28.6943
17.8122 | -35.2923 | -28.7019
17.8183 | -35.2929 | -28.7098
17.8244 | -35.2934 | -28.7174
17.8305 | -35.294 | -28.725
17.8366 | -35.2945 | -28.7324
17.8428 | -35.2951 | -28.7403
17.8489 | -35.2956 | -28.7479
17.855 | -35.2962 | -28.7553
17.8611 | -35.2968 | -28.7628
17.8672 | -35.2973 | -28.7703
17.8733 | -35.2978 | -28.7782
17.8794 | -35.2984 | -28.7859
17.8855 | -35.299 | -28.7935
17.8916 | -35.2995 | -28.8009
17.8977 | -35.3001 | -28.8085
17.9038 | -35.3006 | -28.816
17.9099 | -35.3012 | -28.8237
17.916 | -35.3017 | -28.8312
17.9221 | -35.3023 | -28.8388
17.9282 | -35.3028 | -28.8463
17.9343 | -35.3034 | -28.854
17.9404 | -35.3039 | -28.8615
17.9465 | -35.3045 | -28.8693
17.9526 | -35.305 | -28.8771
17.9587 | -35.3056 | -28.8848
17.9648 | -35.3061 | -28.8924
17.9709 | -35.3067 | -28.8999
17.977 | -35.3073 | -28.9074
17.9832 | -35.3078 | -28.915
17.9893 | -35.3083 | -28.9227
17.9954 | -35.3089 | -28.9304
18.0015 | -35.3095 | -28.938
18.0076 | -35.31 | -28.9452
18.0137 | -35.3106 | -28.9531
18.0198 | -35.3111 | -28.9608
18.0259 | -35.3116 | -28.9684
18.032 | -35.3121 | -28.9757
18.0381 | -35.3127 | -28.9831
18.0442 | -35.3133 | -28.9907
18.0503 | -35.3138 | -28.9984
18.0564 | -35.3144 | -29.0059
18.0625 | -35.3149 | -29.0136
18.0686 | -35.3155 | -29.0212
18.0747 | -35.3161 | -29.0288
18.0808 | -35.3166 | -29.0362
18.0869 | -35.3172 | -29.044
18.093 | -35.3178 | -29.0515
18.0991 | -35.3183 | -29.0592
18.1052 | -35.3189 | -29.0672
18.1113 | -35.3194 | -29.0744
18.1174 | -35.32 | -29.0818
18.1236 | -35.3206 | -29.0895
18.1297 | -35.3211 | -29.0969
18.1358 | -35.3217 | -29.1045
18.1419 | -35.3223 | -29.1121
18.148 | -35.3228 | -29.1198
18.1541 | -35.3235 | -29.1269
18.1602 | -35.324 | -29.1346
18.1663 | -35.3247 | -29.1422
18.1724 | -35.3252 | -29.15
18.1785 | -35.3257 | -29.1577
18.1846 | -35.3262 | -29.165
18.1907 | -35.3268 | -29.1729
18.1968 | -35.3273 | -29.1806
18.2029 | -35.3279 | -29.1878
18.209 | -35.3285 | -29.1956
18.2151 | -35.329 | -29.2032
18.2212 | -35.3296 | -29.2105
18.2273 | -35.3302 | -29.2179
18.2334 | -35.3307 | -29.2255
18.2395 | -35.3312 | -29.2331
18.2456 | -35.3318 | -29.2405
18.2517 | -35.3324 | -29.2479
18.2578 | -35.3329 | -29.2553
18.2639 | -35.3335 | -29.2629
18.2701 | -35.3341 | -29.2705
18.2762 | -35.3346 | -29.2783
18.2823 | -35.3352 | -29.2861
18.2884 | -35.3358 | -29.2935
18.2945 | -35.3364 | -29.3008
18.3006 | -35.337 | -29.3084
18.3067 | -35.3376 | -29.3161
18.3128 | -35.3382 | -29.3235
18.3189 | -35.3388 | -29.3311
18.325 | -35.3393 | -29.3385
18.3311 | -35.3398 | -29.3462
18.3372 | -35.3404 | -29.3536
18.3433 | -35.3409 | -29.3608
18.3494 | -35.3415 | -29.3687
18.3555 | -35.342 | -29.3761
18.3616 | -35.3426 | -29.3835
18.3677 | -35.3432 | -29.3906
18.3738 | -35.3437 | -29.3977
18.3799 | -35.3443 | -29.405
18.386 | -35.3448 | -29.4125
18.3921 | -35.3454 | -29.42
18.3982 | -35.346 | -29.4276
18.4043 | -35.3465 | -29.4351
18.4105 | -35.347 | -29.4426
18.4166 | -35.3476 | -29.4501
18.4227 | -35.3481 | -29.4577
18.4288 | -35.3488 | -29.4654
18.4349 | -35.3493 | -29.4729
18.441 | -35.3499 | -29.4804
18.4471 | -35.3504 | -29.4882
18.4532 | -35.351 | -29.4958
18.4593 | -35.3516 | -29.5033
18.4654 | -35.3521 | -29.5104
18.4715 | -35.3526 | -29.518
18.4776 | -35.3532 | -29.5255
18.4837 | -35.3538 | -29.5332
18.4898 | -35.3544 | -29.5408
18.4959 | -35.355 | -29.5485
18.502 | -35.3556 | -29.556
18.5081 | -35.3562 | -29.5637
18.5142 | -35.3567 | -29.5711
18.5203 | -35.3572 | -29.5784
18.5264 | -35.3577 | -29.5857
18.5325 | -35.3583 | -29.5934
18.5386 | -35.3589 | -29.6011
18.5447 | -35.3595 | -29.6086
18.5508 | -35.3601 | -29.6162
18.557 | -35.3607 | -29.6239
18.5631 | -35.3612 | -29.6314
18.5692 | -35.3617 | -29.6388
18.5753 | -35.3622 | -29.6461
18.5814 | -35.3628 | -29.654
18.5875 | -35.3634 | -29.6618
18.5936 | -35.3639 | -29.6696
18.5997 | -35.3645 | -29.6773
18.6058 | -35.3651 | -29.6847
18.6119 | -35.3656 | -29.6922
18.618 | -35.3662 | -29.6999
18.6241 | -35.3668 | -29.7078
18.6302 | -35.3674 | -29.7152
18.6363 | -35.3679 | -29.7226
18.6424 | -35.3685 | -29.7299
18.6485 | -35.369 | -29.7375
18.6546 | -35.3696 | -29.7447
18.6607 | -35.3702 | -29.7522
18.6668 | -35.3708 | -29.7596
18.6729 | -35.3713 | -29.7671
18.679 | -35.3719 | -29.7744
18.6851 | -35.3725 | -29.7817
18.6912 | -35.3731 | -29.7892
18.6974 | -35.3736 | -29.7966
18.7035 | -35.3742 | -29.8041
18.7096 | -35.3748 | -29.8116
18.7157 | -35.3754 | -29.8187
18.7218 | -35.3759 | -29.8262
18.7279 | -35.3765 | -29.8334
18.734 | -35.3771 | -29.8407
18.7401 | -35.3777 | -29.8479
18.7462 | -35.3782 | -29.8546
18.7523 | -35.3788 | -29.862
18.7584 | -35.3794 | -29.8696
18.7645 | -35.38 | -29.877
18.7706 | -35.3806 | -29.8844
18.7767 | -35.3812 | -29.8919
18.7828 | -35.3818 | -29.8993
18.7889 | -35.3824 | -29.9067
18.795 | -35.3829 | -29.9141
18.8011 | -35.3835 | -29.9214
18.8072 | -35.3841 | -29.9287
18.8133 | -35.3847 | -29.9363
18.8194 | -35.3852 | -29.944
18.8255 | -35.3858 | -29.9511
18.8316 | -35.3863 | -29.9588
18.8377 | -35.3869 | -29.9664
18.8439 | -35.3874 | -29.9738
18.85 | -35.388 | -29.9815
18.8561 | -35.3886 | -29.9884
18.8622 | -35.3892 | -29.996
18.8683 | -35.3898 | -30.003
18.8744 | -35.3904 | -30.0103
18.8805 | -35.391 | -30.0178
18.8866 | -35.3915 | -30.0255
18.8927 | -35.3922 | -30.0324
18.8988 | -35.3927 | -30.0398
18.9049 | -35.3933 | -30.0471
18.911 | -35.3939 | -30.0543
18.9171 | -35.3945 | -30.0617
18.9232 | -35.3951 | -30.0689
18.9293 | -35.3956 | -30.0762
18.9354 | -35.3961 | -30.0832
18.9415 | -35.3967 | -30.0904
18.9476 | -35.3972 | -30.0979
18.9537 | -35.3978 | -30.1053
18.9598 | -35.3984 | -30.1122
18.9659 | -35.399 | -30.1195
18.972 | -35.3996 | -30.1268
18.9781 | -35.4001 | -30.1341
18.9843 | -35.4007 | -30.1416
18.9904 | -35.4012 | -30.149
18.9965 | -35.4018 | -30.1561
19.0026 | -35.4024 | -30.1636
19.0087 | -35.403 | -30.1709
19.0148 | -35.4036 | -30.1782
19.0209 | -35.4041 | -30.186
19.027 | -35.4047 | -30.1935
19.0331 | -35.4053 | -30.2009
19.0392 | -35.406 | -30.2082
19.0453 | -35.4065 | -30.2155
19.0514 | -35.407 | -30.223
19.0575 | -35.4076 | -30.2304
19.0636 | -35.4081 | -30.2372
19.0697 | -35.4087 | -30.245
19.0758 | -35.4093 | -30.2521
19.0819 | -35.4099 | -30.2598
19.088 | -35.4104 | -30.2671
19.0941 | -35.4109 | -30.274
19.1002 | -35.4115 | -30.2809
19.1063 | -35.4121 | -30.2884
19.1124 | -35.4127 | -30.2956
19.1185 | -35.4133 | -30.3034
19.1246 | -35.4138 | -30.3106
19.1308 | -35.4144 | -30.317
19.1369 | -35.4151 | -30.3241
19.143 | -35.4156 | -30.3314
19.1491 | -35.4162 | -30.3386
19.1552 | -35.4168 | -30.3462
19.1613 | -35.4174 | -30.3535
19.1674 | -35.4179 | -30.361
19.1735 | -35.4184 | -30.3681
19.1796 | -35.419 | -30.3753
19.1857 | -35.4196 | -30.3828
19.1918 | -35.4202 | -30.3903
19.1979 | -35.4208 | -30.3974
19.204 | -35.4214 | -30.4046
19.2101 | -35.422 | -30.4119
19.2162 | -35.4225 | -30.4194
19.2223 | -35.4231 | -30.4262
19.2284 | -35.4236 | -30.4339
19.2345 | -35.4243 | -30.4414
19.2406 | -35.4248 | -30.4486
19.2467 | -35.4253 | -30.4559
19.2528 | -35.4259 | -30.4627
19.2589 | -35.4264 | -30.47
19.265 | -35.4271 | -30.4771
19.2712 | -35.4276 | -30.4843
19.2773 | -35.4282 | -30.491
19.2834 | -35.4287 | -30.498
19.2895 | -35.4294 | -30.505
19.2956 | -35.4299 | -30.5125
19.3017 | -35.4304 | -30.5192
19.3078 | -35.431 | -30.5271
19.3139 | -35.4316 | -30.5338
19.32 | -35.4322 | -30.5412
19.3261 | -35.4326 | -30.5487
19.3322 | -35.4332 | -30.5558
19.3383 | -35.4338 | -30.5625
19.3444 | -35.4344 | -30.5692
19.3505 | -35.435 | -30.5764
19.3566 | -35.4356 | -30.5834
19.3627 | -35.4362 | -30.591
19.3688 | -35.4368 | -30.5986
19.3749 | -35.4373 | -30.6058
19.381 | -35.4379 | -30.6125
19.3871 | -35.4385 | -30.6199
19.3932 | -35.4391 | -30.627
19.3993 | -35.4397 | -30.6339
19.4054 | -35.4403 | -30.6412
19.4115 | -35.4409 | -30.6482
19.4177 | -35.4415 | -30.6551
19.4238 | -35.4421 | -30.662
19.4299 | -35.4427 | -30.6689
19.436 | -35.4433 | -30.6765
19.4421 | -35.4439 | -30.6831
19.4482 | -35.4445 | -30.6901
19.4543 | -35.4451 | -30.6968
19.4604 | -35.4456 | -30.7038
19.4665 | -35.4462 | -30.711
19.4726 | -35.4468 | -30.718
19.4787 | -35.4474 | -30.7254
19.4848 | -35.448 | -30.7325
19.4909 | -35.4485 | -30.7394
19.497 | -35.4491 | -30.7466
19.5031 | -35.4496 | -30.7539
19.5092 | -35.4502 | -30.7611
19.5153 | -35.4508 | -30.7684
19.5214 | -35.4514 | -30.7753
19.5275 | -35.4519 | -30.7821
19.5336 | -35.4525 | -30.789
19.5397 | -35.4531 | -30.7962
19.5458 | -35.4537 | -30.8035
19.5519 | -35.4543 | -30.8103
19.5581 | -35.4549 | -30.8169
19.5642 | -35.4554 | -30.8238
19.5703 | -35.456 | -30.8307
19.5764 | -35.4566 | -30.8378
19.5825 | -35.4572 | -30.8449
19.5886 | -35.4577 | -30.852
19.5947 | -35.4583 | -30.859
19.6008 | -35.4588 | -30.866
19.6069 | -35.4594 | -30.8729
19.613 | -35.4599 | -30.88
19.6191 | -35.4605 | -30.8872
19.6252 | -35.4611 | -30.8943
19.6313 | -35.4616 | -30.9017
19.6374 | -35.4623 | -30.9086
19.6435 | -35.4629 | -30.9157
19.6496 | -35.4635 | -30.9225
19.6557 | -35.464 | -30.9295
19.6618 | -35.4646 | -30.9364
19.6679 | -35.4652 | -30.9436
19.674 | -35.4658 | -30.9507
19.6801 | -35.4663 | -30.9576
19.6862 | -35.4669 | -30.9642
19.6923 | -35.4674 | -30.9711
19.6984 | -35.468 | -30.9775
19.7046 | -35.4686 | -30.9846
19.7107 | -35.4692 | -30.9913
19.7168 | -35.4698 | -30.9989
19.7229 | -35.4703 | -31.0064
19.729 | -35.4708 | -31.0138
19.7351 | -35.4714 | -31.0202
19.7412 | -35.472 | -31.0271
19.7473 | -35.4725 | -31.0347
19.7534 | -35.4731 | -31.0417
19.7595 | -35.4736 | -31.0488
19.7656 | -35.4742 | -31.0561
19.7717 | -35.4748 | -31.0624
19.7778 | -35.4753 | -31.0697
19.7839 | -35.4759 | -31.0761
19.79 | -35.4765 | -31.0833
19.7961 | -35.477 | -31.0904
19.8022 | -35.4776 | -31.0975
19.8083 | -35.4782 | -31.1046
19.8144 | -35.4787 | -31.1118
19.8205 | -35.4793 | -31.1191
19.8266 | -35.4799 | -31.1261
19.8327 | -35.4804 | -31.1333
19.8388 | -35.481 | -31.1406
19.845 | -35.4816 | -31.1474
19.8511 | -35.4822 | -31.1541
19.8572 | -35.4828 | -31.1613
19.8633 | -35.4834 | -31.1684
19.8694 | -35.484 | -31.1757
19.8755 | -35.4847 | -31.1828
19.8816 | -35.4853 | -31.1905
19.8877 | -35.4859 | -31.1972
19.8938 | -35.4864 | -31.204
19.8999 | -35.487 | -31.2107
19.906 | -35.4877 | -31.2177
19.9121 | -35.4883 | -31.2244
19.9182 | -35.4889 | -31.2313
19.9243 | -35.4894 | -31.2384
19.9304 | -35.49 | -31.2458
19.9365 | -35.4905 | -31.2528
19.9426 | -35.491 | -31.2595
19.9487 | -35.4916 | -31.2661
19.9548 | -35.4922 | -31.2734
19.9609 | -35.4928 | -31.2801
19.967 | -35.4934 | -31.2872
19.9731 | -35.494 | -31.2942
19.9792 | -35.4946 | -31.3011
19.9853 | -35.4952 | -31.3081
19.9915 | -35.4958 | -31.3153
19.9976 | -35.4964 | -31.3223
20.0037 | -35.497 | -31.3292
20.0098 | -35.4974 | -31.3361
20.0159 | -35.498 | -31.3427
20.022 | -35.4987 | -31.35
20.0281 | -35.4993 | -31.3567
20.0342 | -35.4998 | -31.3633
20.0403 | -35.5005 | -31.3699
20.0464 | -35.501 | -31.3775
20.0525 | -35.5016 | -31.3844
20.0586 | -35.5022 | -31.3908
20.0647 | -35.5027 | -31.3982
20.0708 | -35.5034 | -31.4054
20.0769 | -35.504 | -31.412
20.083 | -35.5045 | -31.4186
20.0891 | -35.505 | -31.4256
20.0952 | -35.5056 | -31.432
20.1013 | -35.5062 | -31.4388
20.1074 | -35.5068 | -31.4458
20.1135 | -35.5073 | -31.4528
20.1196 | -35.5079 | -31.4595
20.1257 | -35.5085 | -31.4662
20.1319 | -35.509 | -31.4729
20.138 | -35.5096 | -31.4801
20.1441 | -35.5102 | -31.4869
20.1502 | -35.5108 | -31.494
20.1563 | -35.5114 | -31.501
20.1624 | -35.5119 | -31.5082
20.1685 | -35.5125 | -31.515
20.1746 | -35.5131 | -31.5221
20.1807 | -35.5137 | -31.5288
20.1868 | -35.5142 | -31.5359
20.1929 | -35.5147 | -31.5425
20.199 | -35.5153 | -31.5497
20.2051 | -35.5158 | -31.5565
20.2112 | -35.5164 | -31.5633
20.2173 | -35.5169 | -31.5701
20.2234 | -35.5175 | -31.5772
20.2295 | -35.5181 | -31.5834
20.2356 | -35.5187 | -31.5901
20.2417 | -35.5193 | -31.5969
20.2478 | -35.5198 | -31.6036
20.2539 | -35.5203 | -31.6108
20.26 | -35.5209 | -31.6173
20.2661 | -35.5215 | -31.6242
20.2723 | -35.522 | -31.6305
20.2784 | -35.5226 | -31.638
20.2845 | -35.5231 | -31.645
20.2906 | -35.5237 | -31.6521
20.2967 | -35.5244 | -31.6588
20.3028 | -35.525 | -31.6655
20.3089 | -35.5255 | -31.6726
20.315 | -35.5261 | -31.679
20.3211 | -35.5266 | -31.6859
20.3272 | -35.5272 | -31.6925
20.3333 | -35.5278 | -31.6992
20.3394 | -35.5283 | -31.7057
20.3455 | -35.5289 | -31.7124
20.3516 | -35.5294 | -31.7197
20.3577 | -35.53 | -31.7264
20.3638 | -35.5306 | -31.7332
20.3699 | -35.5312 | -31.7399
20.376 | -35.5318 | -31.7469
20.3821 | -35.5324 | -31.7533
20.3882 | -35.5329 | -31.7599
20.3943 | -35.5335 | -31.767
20.4004 | -35.534 | -31.7732
20.4065 | -35.5346 | -31.7803
20.4126 | -35.5352 | -31.787
20.4188 | -35.5358 | -31.7942
20.4249 | -35.5364 | -31.8009
20.431 | -35.5369 | -31.8083
20.4371 | -35.5375 | -31.8142
20.4432 | -35.538 | -31.8212
20.4493 | -35.5386 | -31.8283
20.4554 | -35.5392 | -31.8352
20.4615 | -35.5398 | -31.842
20.4676 | -35.5404 | -31.8485
20.4737 | -35.5409 | -31.8549
20.4798 | -35.5415 | -31.8617
20.4859 | -35.5421 | -31.8681
20.492 | -35.5426 | -31.8745
20.4981 | -35.5432 | -31.8814
20.5042 | -35.5438 | -31.8878
20.5103 | -35.5444 | -31.8942
20.5164 | -35.5449 | -31.9007
20.5225 | -35.5454 | -31.9075
20.5286 | -35.5459 | -31.9148
20.5347 | -35.5465 | -31.9214
20.5408 | -35.5471 | -31.9283
20.5469 | -35.5477 | -31.9351
20.553 | -35.5483 | -31.942
20.5592 | -35.5489 | -31.9483
20.5653 | -35.5494 | -31.955
20.5714 | -35.5499 | -31.9616
20.5775 | -35.5504 | -31.9679
20.5836 | -35.5509 | -31.9747
20.5897 | -35.5514 | -31.9812
20.5958 | -35.5519 | -31.9879
20.6019 | -35.5525 | -31.9947
20.608 | -35.553 | -32.0015
20.6141 | -35.5535 | -32.0076
20.6202 | -35.5541 | -32.0145
20.6263 | -35.5546 | -32.0207
20.6324 | -35.5552 | -32.0271
20.6385 | -35.5557 | -32.0338
20.6446 | -35.5563 | -32.0411
20.6507 | -35.5568 | -32.0478
20.6568 | -35.5573 | -32.0545
20.6629 | -35.5579 | -32.0606
20.669 | -35.5584 | -32.0675
20.6751 | -35.559 | -32.0739
20.6812 | -35.5596 | -32.0799
20.6873 | -35.5602 | -32.0867
20.6934 | -35.5608 | -32.0934
20.6995 | -35.5613 | -32.1002
20.7057 | -35.5619 | -32.107
20.7118 | -35.5625 | -32.1134
20.7179 | -35.5632 | -32.1198
20.724 | -35.5637 | -32.1261
20.7301 | -35.5642 | -32.133
20.7362 | -35.5648 | -32.1398
20.7423 | -35.5653 | -32.1463
20.7484 | -35.5659 | -32.153
20.7545 | -35.5664 | -32.1596
20.7606 | -35.567 | -32.1664
20.7667 | -35.5677 | -32.1735
20.7728 | -35.5683 | -32.1799
20.7789 | -35.5688 | -32.1867
20.785 | -35.5693 | -32.1932
20.7911 | -35.5699 | -32.1992
20.7972 | -35.5704 | -32.2058
20.8033 | -35.5709 | -32.2126
20.8094 | -35.5715 | -32.2197
20.8155 | -35.572 | -32.2263
20.8216 | -35.5725 | -32.2332
20.8277 | -35.5731 | -32.2396
20.8338 | -35.5736 | -32.246
20.8399 | -35.5742 | -32.2522
20.8461 | -35.5747 | -32.2583
20.8522 | -35.5752 | -32.2647
20.8583 | -35.5758 | -32.2708
20.8644 | -35.5764 | -32.2774
20.8705 | -35.577 | -32.2842
20.8766 | -35.5775 | -32.2912
20.8827 | -35.5781 | -32.2975
20.8888 | -35.5785 | -32.3038
20.8949 | -35.579 | -32.3103
20.901 | -35.5795 | -32.3172
20.9071 | -35.5801 | -32.3236
20.9132 | -35.5806 | -32.3303
20.9193 | -35.5812 | -32.3368
20.9254 | -35.5817 | -32.3429
20.9315 | -35.5823 | -32.3491
20.9376 | -35.5828 | -32.3558
20.9437 | -35.5834 | -32.3627
20.9498 | -35.5839 | -32.3693
20.9559 | -35.5845 | -32.3754
20.962 | -35.5851 | -32.3819
20.9681 | -35.5856 | -32.3885
20.9742 | -35.5862 | -32.395
20.9803 | -35.5867 | -32.4016
20.9864 | -35.5873 | -32.408
20.9926 | -35.5878 | -32.4146
20.9987 | -35.5883 | -32.421
21.0048 | -35.5889 | -32.4272
21.0109 | -35.5895 | -32.4336
21.017 | -35.59 | -32.4397
21.0231 | -35.5906 | -32.4459
21.0292 | -35.5911 | -32.4523
21.0353 | -35.5916 | -32.4594
21.0414 | -35.5922 | -32.4658
21.0475 | -35.5928 | -32.4728
21.0536 | -35.5933 | -32.4794
21.0597 | -35.5938 | -32.4859
21.0658 | -35.5943 | -32.4927
21.0719 | -35.5949 | -32.4993
21.078 | -35.5954 | -32.5059
21.0841 | -35.596 | -32.5124
21.0902 | -35.5966 | -32.519
21.0963 | -35.5971 | -32.5256
21.1024 | -35.5976 | -32.5319
21.1085 | -35.5981 | -32.538
21.1146 | -35.5986 | -32.5447
21.1207 | -35.5992 | -32.5508
21.1268 | -35.5997 | -32.5573
21.133 | -35.6002 | -32.564
21.1391 | -35.6007 | -32.5697
21.1452 | -35.6013 | -32.5764
21.1513 | -35.6019 | -32.5831
21.1574 | -35.6024 | -32.5896
21.1635 | -35.603 | -32.5961
21.1696 | -35.6035 | -32.6027
21.1757 | -35.604 | -32.6091
21.1818 | -35.6046 | -32.6156
21.1879 | -35.6051 | -32.6224
21.194 | -35.6057 | -32.6292
21.2001 | -35.6062 | -32.6357
21.2062 | -35.6067 | -32.6423
21.2123 | -35.6072 | -32.6483
21.2184 | -35.6078 | -32.6546
21.2245 | -35.6084 | -32.6612
21.2306 | -35.609 | -32.6679
21.2367 | -35.6095 | -32.675
21.2428 | -35.61 | -32.6817
21.2489 | -35.6105 | -32.6876
21.255 | -35.611 | -32.6944
21.2611 | -35.6116 | -32.7009
21.2672 | -35.6121 | -32.7073
21.2733 | -35.6126 | -32.7131
21.2795 | -35.6132 | -32.719
21.2856 | -35.6137 | -32.7257
21.2917 | -35.6143 | -32.7317
21.2978 | -35.6148 | -32.7378
21.3039 | -35.6153 | -32.7446
21.31 | -35.6158 | -32.7506
21.3161 | -35.6163 | -32.7568
21.3222 | -35.6168 | -32.7638
21.3283 | -35.6174 | -32.7705
21.3344 | -35.6179 | -32.7772
21.3405 | -35.6185 | -32.7836
21.3466 | -35.619 | -32.7898
21.3527 | -35.6196 | -32.7958
21.3588 | -35.6201 | -32.8024
21.3649 | -35.6207 | -32.8086
21.371 | -35.6213 | -32.8149
21.3771 | -35.6218 | -32.8212
21.3832 | -35.6223 | -32.8278
21.3893 | -35.6227 | -32.8334
21.3954 | -35.6233 | -32.8398
21.4015 | -35.6238 | -32.8462
21.4076 | -35.6243 | -32.8526
21.4137 | -35.6249 | -32.8593
21.4199 | -35.6254 | -32.8658
21.426 | -35.6259 | -32.8726
21.4321 | -35.6264 | -32.8787
21.4382 | -35.6269 | -32.885
21.4443 | -35.6274 | -32.8915
21.4504 | -35.628 | -32.8978
21.4565 | -35.6285 | -32.9039
21.4626 | -35.6291 | -32.9102
21.4687 | -35.6296 | -32.9167
21.4748 | -35.6301 | -32.9225
21.4809 | -35.6306 | -32.929
21.487 | -35.6311 | -32.9351
21.4931 | -35.6317 | -32.9411
21.4992 | -35.6322 | -32.9475
21.5053 | -35.6327 | -32.9538
21.5114 | -35.6333 | -32.96
21.5175 | -35.6337 | -32.9662
21.5236 | -35.6342 | -32.9724
21.5297 | -35.6348 | -32.9788
21.5358 | -35.6353 | -32.9849
21.5419 | -35.6358 | -32.9918
21.548 | -35.6364 | -32.9981
21.5541 | -35.6369 | -33.0045
21.5602 | -35.6375 | -33.011
21.5664 | -35.638 | -33.0173
21.5725 | -35.6384 | -33.0239
21.5786 | -35.6389 | -33.03
21.5847 | -35.6394 | -33.0359
21.5908 | -35.64 | -33.0424
21.5969 | -35.6405 | -33.0477
21.603 | -35.6411 | -33.0543
21.6091 | -35.6416 | -33.0605
21.6152 | -35.6422 | -33.0671
21.6213 | -35.6427 | -33.0739
21.6274 | -35.6433 | -33.0803
21.6335 | -35.6438 | -33.0868
21.6396 | -35.6444 | -33.093
21.6457 | -35.6449 | -33.0994
21.6518 | -35.6454 | -33.1055
21.6579 | -35.6459 | -33.1117
21.664 | -35.6465 | -33.1181
21.6701 | -35.6471 | -33.1241
21.6762 | -35.6476 | -33.1304
21.6823 | -35.6481 | -33.1367
21.6884 | -35.6486 | -33.1434
21.6945 | -35.6492 | -33.1496
21.7006 | -35.6497 | -33.1559
21.7068 | -35.6502 | -33.1618
21.7129 | -35.6507 | -33.1687
21.719 | -35.6512 | -33.1751
21.7251 | -35.6518 | -33.1816
21.7312 | -35.6523 | -33.1876
21.7373 | -35.6528 | -33.1941
21.7434 | -35.6534 | -33.2003
21.7495 | -35.6539 | -33.2065
21.7556 | -35.6545 | -33.2132
21.7617 | -35.655 | -33.219
21.7678 | -35.6556 | -33.2254
21.7739 | -35.6561 | -33.2316
21.78 | -35.6567 | -33.2381
21.7861 | -35.6573 | -33.2448
21.7922 | -35.6578 | -33.2511
21.7983 | -35.6584 | -33.257
21.8044 | -35.6589 | -33.2636
21.8105 | -35.6594 | -33.2693
21.8166 | -35.66 | -33.2758
21.8227 | -35.6605 | -33.2822
21.8288 | -35.661 | -33.2882
21.8349 | -35.6615 | -33.2943
21.841 | -35.6621 | -33.3008
21.8471 | -35.6626 | -33.3073
21.8533 | -35.6631 | -33.3135
21.8594 | -35.6636 | -33.3199
21.8655 | -35.6642 | -33.3267
21.8716 | -35.6647 | -33.3332
21.8777 | -35.6652 | -33.3392
21.8838 | -35.6657 | -33.3455
21.8899 | -35.6662 | -33.3515
21.896 | -35.6667 | -33.3578
21.9021 | -35.6672 | -33.3644
21.9082 | -35.6676 | -33.3707
21.9143 | -35.6682 | -33.3762
21.9204 | -35.6687 | -33.3823
21.9265 | -35.6692 | -33.3886
21.9326 | -35.6697 | -33.3952
21.9387 | -35.6703 | -33.4013
21.9448 | -35.6709 | -33.4077
21.9509 | -35.6714 | -33.4135
21.957 | -35.6719 | -33.4199
21.9631 | -35.6724 | -33.4263
21.9692 | -35.673 | -33.4328
21.9753 | -35.6735 | -33.4385
21.9814 | -35.6741 | -33.4447
21.9875 | -35.6745 | -33.4506
21.9937 | -35.675 | -33.4573
21.9998 | -35.6755 | -33.4636
22.0059 | -35.6762 | -33.4698
22.012 | -35.6768 | -33.4762
22.0181 | -35.6773 | -33.4824
22.0242 | -35.6779 | -33.4889
22.0303 | -35.6784 | -33.4951
22.0364 | -35.6788 | -33.5017
22.0425 | -35.6793 | -33.5083
22.0486 | -35.6797 | -33.5146
22.0547 | -35.6802 | -33.5208
22.0608 | -35.6808 | -33.527
22.0669 | -35.6813 | -33.533
22.073 | -35.6818 | -33.5391
22.0791 | -35.6824 | -33.5452
22.0852 | -35.6829 | -33.5515
22.0913 | -35.6834 | -33.5579
22.0974 | -35.6839 | -33.5639
22.1035 | -35.6844 | -33.5698
22.1096 | -35.6849 | -33.5764
22.1157 | -35.6854 | -33.5828
22.1218 | -35.686 | -33.5893
22.1279 | -35.6865 | -33.5958
22.134 | -35.687 | -33.6018
22.1402 | -35.6875 | -33.6077
22.1463 | -35.6881 | -33.6144
22.1524 | -35.6885 | -33.6205
22.1585 | -35.6891 | -33.6269
22.1646 | -35.6895 | -33.6329
22.1707 | -35.69 | -33.6397
22.1768 | -35.6906 | -33.646
22.1829 | -35.6911 | -33.6524
22.189 | -35.6916 | -33.6591
22.1951 | -35.6921 | -33.6651
22.2012 | -35.6926 | -33.6715
22.2073 | -35.6931 | -33.6779
22.2134 | -35.6936 | -33.6841
22.2195 | -35.6942 | -33.6901
22.2256 | -35.6947 | -33.696
22.2317 | -35.6952 | -33.7018
22.2378 | -35.6957 | -33.708
22.2439 | -35.6963 | -33.7143
22.25 | -35.6969 | -33.7207
22.2561 | -35.6975 | -33.7269
22.2622 | -35.6979 | -33.7333
22.2683 | -35.6985 | -33.7393
22.2744 | -35.699 | -33.7453
22.2806 | -35.6995 | -33.7515
22.2867 | -35.7 | -33.7578
22.2928 | -35.7005 | -33.7641
22.2989 | -35.701 | -33.7701
22.305 | -35.7016 | -33.7764
22.3111 | -35.702 | -33.7828
22.3172 | -35.7026 | -33.7893
22.3233 | -35.7031 | -33.7953
22.3294 | -35.7036 | -33.8015
22.3355 | -35.7042 | -33.8075
22.3416 | -35.7047 | -33.8141
22.3477 | -35.7052 | -33.8203
22.3538 | -35.7058 | -33.8267
22.3599 | -35.7062 | -33.8325
22.366 | -35.7068 | -33.8386
22.3721 | -35.7072 | -33.8447
22.3782 | -35.7078 | -33.8512
22.3843 | -35.7083 | -33.8571
22.3904 | -35.7088 | -33.8634
22.3965 | -35.7093 | -33.8697
22.4026 | -35.7098 | -33.8755
22.4087 | -35.7103 | -33.8814
22.4148 | -35.7108 | -33.8876
22.4209 | -35.7113 | -33.8935
22.4271 | -35.7118 | -33.8998
22.4332 | -35.7123 | -33.9056
22.4393 | -35.7129 | -33.9117
22.4454 | -35.7133 | -33.9175
22.4515 | -35.7137 | -33.9239
22.4576 | -35.7143 | -33.93
22.4637 | -35.7148 | -33.9369
22.4698 | -35.7153 | -33.9429
22.4759 | -35.7158 | -33.9493
22.482 | -35.7163 | -33.9556
22.4881 | -35.7168 | -33.9618
22.4942 | -35.7173 | -33.9682
22.5003 | -35.7178 | -33.9744
22.5064 | -35.7183 | -33.9807
22.5125 | -35.7188 | -33.9868
22.5186 | -35.7193 | -33.9927
22.5247 | -35.7198 | -33.9989
22.5308 | -35.7203 | -34.0053
22.5369 | -35.7209 | -34.0116
22.543 | -35.7214 | -34.0181
22.5491 | -35.7219 | -34.0243
22.5552 | -35.7224 | -34.0304
22.5613 | -35.7229 | -34.0365
22.5675 | -35.7234 | -34.0429
22.5736 | -35.7239 | -34.0495
22.5797 | -35.7243 | -34.0553
22.5858 | -35.7248 | -34.0608
22.5919 | -35.7253 | -34.0674
22.598 | -35.7258 | -34.0737
22.6041 | -35.7263 | -34.08
22.6102 | -35.7268 | -34.0863
22.6163 | -35.7272 | -34.0923
22.6224 | -35.7277 | -34.0986
22.6285 | -35.7282 | -34.1043
22.6346 | -35.7287 | -34.1103
22.6407 | -35.7292 | -34.1164
22.6468 | -35.7297 | -34.1222
22.6529 | -35.7302 | -34.1281
22.659 | -35.7307 | -34.1341
22.6651 | -35.7312 | -34.1401
22.6712 | -35.7317 | -34.1464
22.6773 | -35.7321 | -34.1525
22.6834 | -35.7326 | -34.1587
22.6895 | -35.7331 | -34.1647
22.6956 | -35.7337 | -34.1709
22.7017 | -35.7341 | -34.1774
22.7079 | -35.7346 | -34.1837
22.714 | -35.7351 | -34.1898
22.7201 | -35.7356 | -34.1958
22.7262 | -35.7361 | -34.2021
22.7323 | -35.7366 | -34.2083
22.7384 | -35.7371 | -34.2143
22.7445 | -35.7376 | -34.2201
22.7506 | -35.7381 | -34.2264
22.7567 | -35.7386 | -34.2327
22.7628 | -35.7391 | -34.239
22.7689 | -35.7396 | -34.245
22.775 | -35.7401 | -34.2511
22.7811 | -35.7406 | -34.2573
22.7872 | -35.741 | -34.2637
22.7933 | -35.7415 | -34.2694
22.7994 | -35.7421 | -34.2754
22.8055 | -35.7426 | -34.2816
22.8116 | -35.7431 | -34.2878
22.8177 | -35.7436 | -34.2939
22.8238 | -35.7441 | -34.3001
22.8299 | -35.7446 | -34.3059
22.836 | -35.7451 | -34.3123
22.8421 | -35.7456 | -34.3185
22.8482 | -35.7461 | -34.3245
22.8544 | -35.7466 | -34.3308
22.8605 | -35.7472 | -34.3371
22.8666 | -35.7476 | -34.3434
22.8727 | -35.7481 | -34.3492
22.8788 | -35.7486 | -34.355
22.8849 | -35.749 | -34.361
22.891 | -35.7495 | -34.3674
22.8971 | -35.75 | -34.3735
22.9032 | -35.7505 | -34.3795
22.9093 | -35.751 | -34.3855
22.9154 | -35.7515 | -34.392
22.9215 | -35.752 | -34.3981
22.9276 | -35.7524 | -34.4039
22.9337 | -35.753 | -34.4102
22.9398 | -35.7534 | -34.4163
22.9459 | -35.7539 | -34.4225
22.952 | -35.7544 | -34.4286
22.9581 | -35.7549 | -34.4349
22.9642 | -35.7554 | -34.4412
22.9703 | -35.7559 | -34.4468
22.9764 | -35.7564 | -34.4529
22.9825 | -35.7569 | -34.4591
22.9886 | -35.7574 | -34.465
22.9948 | -35.7578 | -34.4711
23.0009 | -35.7583 | -34.477
23.007 | -35.7588 | -34.4829
23.0131 | -35.7592 | -34.4889
23.0192 | -35.7597 | -34.4948
23.0253 | -35.7602 | -34.5007
23.0314 | -35.7607 | -34.507
23.0375 | -35.7612 | -34.5131
23.0436 | -35.7617 | -34.5186
23.0497 | -35.7623 | -34.5246
23.0558 | -35.7628 | -34.5308
23.0619 | -35.7633 | -34.5371
23.068 | -35.764 | -34.5433
23.0741 | -35.7645 | -34.5492
23.0802 | -35.765 | -34.5555
23.0863 | -35.7655 | -34.5618
23.0924 | -35.766 | -34.5679
23.0985 | -35.7665 | -34.5742
23.1046 | -35.767 | -34.5805
23.1107 | -35.7675 | -34.5862
23.1168 | -35.768 | -34.5925
23.1229 | -35.7684 | -34.5986
23.129 | -35.7689 | -34.6048
23.1351 | -35.7694 | -34.611
23.1413 | -35.7699 | -34.6171
23.1474 | -35.7705 | -34.6226
23.1535 | -35.7711 | -34.6288
23.1596 | -35.7715 | -34.6349
23.1657 | -35.772 | -34.641
23.1718 | -35.7725 | -34.6473
23.1779 | -35.773 | -34.6536
23.184 | -35.7735 | -34.6591
23.1901 | -35.774 | -34.6652
23.1962 | -35.7745 | -34.6711
23.2023 | -35.775 | -34.6774
23.2084 | -35.7755 | -34.6835
23.2145 | -35.776 | -34.6893
23.2206 | -35.7766 | -34.6955
23.2267 | -35.777 | -34.7017
23.2328 | -35.7775 | -34.7077
23.2389 | -35.778 | -34.7139
23.245 | -35.7785 | -34.7201
23.2511 | -35.779 | -34.7261
23.2572 | -35.7796 | -34.7319
23.2633 | -35.7801 | -34.7382
23.2694 | -35.7806 | -34.7443
23.2755 | -35.7811 | -34.7505
23.2817 | -35.7815 | -34.7567
23.2878 | -35.782 | -34.7629
23.2939 | -35.7825 | -34.7694
23.3 | -35.7831 | -34.776
23.3061 | -35.7835 | -34.7823
23.3122 | -35.784 | -34.7885
23.3183 | -35.7844 | -34.7949
23.3244 | -35.785 | -34.8002
23.3305 | -35.7854 | -34.8067
23.3366 | -35.7859 | -34.8131
23.3427 | -35.7864 | -34.819
23.3488 | -35.7868 | -34.8253
23.3549 | -35.7874 | -34.8316
23.361 | -35.7879 | -34.8375
23.3671 | -35.7884 | -34.8435
23.3732 | -35.7889 | -34.8497
23.3793 | -35.7895 | -34.8559
23.3854 | -35.79 | -34.8619
23.3915 | -35.7904 | -34.868
23.3976 | -35.7909 | -34.8742
23.4037 | -35.7914 | -34.88
23.4098 | -35.7919 | -34.8863
23.4159 | -35.7924 | -34.8924
23.422 | -35.7929 | -34.8982
23.4282 | -35.7934 | -34.9041
23.4343 | -35.7939 | -34.9099
23.4404 | -35.7944 | -34.9164
23.4465 | -35.7949 | -34.9227
23.4526 | -35.7954 | -34.9289
23.4587 | -35.7959 | -34.9347
23.4648 | -35.7964 | -34.9404
23.4709 | -35.7969 | -34.9462
23.477 | -35.7974 | -34.9523
23.4831 | -35.798 | -34.9586
23.4892 | -35.7985 | -34.9649
23.4953 | -35.799 | -34.9712
23.5014 | -35.7995 | -34.9775
23.5075 | -35.8 | -34.9836
23.5136 | -35.8005 | -34.9894
23.5197 | -35.801 | -34.9955
23.5258 | -35.8016 | -35.0015
23.5319 | -35.802 | -35.0071
23.538 | -35.8025 | -35.0131
23.5441 | -35.803 | -35.0188
23.5502 | -35.8035 | -35.0251
23.5563 | -35.804 | -35.0311
23.5624 | -35.8045 | -35.037
23.5686 | -35.8049 | -35.0436
23.5747 | -35.8053 | -35.0498
23.5808 | -35.8058 | -35.0558
23.5869 | -35.8063 | -35.0617
23.593 | -35.8068 | -35.0675
23.5991 | -35.8073 | -35.0734
23.6052 | -35.8078 | -35.0788
23.6113 | -35.8083 | -35.0847
23.6174 | -35.8087 | -35.0905
23.6235 | -35.8092 | -35.0963
23.6296 | -35.8097 | -35.1029
23.6357 | -35.8101 | -35.1093
23.6418 | -35.8106 | -35.1151
23.6479 | -35.811 | -35.1218
23.654 | -35.8115 | -35.1277
23.6601 | -35.8119 | -35.1336
23.6662 | -35.8123 | -35.14
23.6723 | -35.8129 | -35.1461
23.6784 | -35.8134 | -35.152
23.6845 | -35.814 | -35.1581
23.6906 | -35.8145 | -35.1642
23.6967 | -35.8149 | -35.1701
23.7028 | -35.8154 | -35.1763
23.7089 | -35.8159 | -35.1825
23.7151 | -35.8164 | -35.1883
23.7212 | -35.8169 | -35.1943
23.7273 | -35.8175 | -35.2005
23.7334 | -35.818 | -35.2064
23.7395 | -35.8185 | -35.2117
23.7456 | -35.8189 | -35.2179
23.7517 | -35.8194 | -35.224
23.7578 | -35.8198 | -35.2299
23.7639 | -35.8204 | -35.2358
23.77 | -35.8209 | -35.2417
23.7761 | -35.8214 | -35.2479
23.7822 | -35.822 | -35.2538
23.7883 | -35.8224 | -35.2597
23.7944 | -35.8229 | -35.2659
23.8005 | -35.8234 | -35.2721
23.8066 | -35.824 | -35.278
23.8127 | -35.8244 | -35.2847
23.8188 | -35.8249 | -35.2907
23.8249 | -35.8254 | -35.2969
23.831 | -35.8259 | -35.3031
23.8371 | -35.8263 | -35.309
23.8432 | -35.8268 | -35.3151
23.8493 | -35.8274 | -35.3209
23.8555 | -35.8279 | -35.327
23.8616 | -35.8284 | -35.3331
23.8677 | -35.8289 | -35.3389
23.8738 | -35.8294 | -35.3452
23.8799 | -35.8299 | -35.3514
23.886 | -35.8303 | -35.3573
23.8921 | -35.8308 | -35.3631
23.8982 | -35.8313 | -35.3692
23.9043 | -35.8318 | -35.3752
23.9104 | -35.8323 | -35.3811
23.9165 | -35.8327 | -35.3869
23.9226 | -35.8332 | -35.3928
23.9287 | -35.8337 | -35.3987
23.9348 | -35.8342 | -35.4044
23.9409 | -35.8347 | -35.4106
23.947 | -35.8351 | -35.4168
23.9531 | -35.8355 | -35.4224
23.9592 | -35.836 | -35.4289
23.9653 | -35.8365 | -35.4351
23.9714 | -35.837 | -35.4407
23.9775 | -35.8374 | -35.4469
23.9836 | -35.838 | -35.4533
23.9897 | -35.8386 | -35.4588
23.9958 | -35.8391 | -35.4652
24.002 | -35.8396 | -35.4713
24.0081 | -35.84 | -35.4774
24.0142 | -35.8405 | -35.4833
24.0203 | -35.841 | -35.4891
24.0264 | -35.8414 | -35.4951
24.0325 | -35.8419 | -35.5014
24.0386 | -35.8423 | -35.5076
24.0447 | -35.8428 | -35.5132
24.0508 | -35.8433 | -35.5195
24.0569 | -35.8438 | -35.5255
24.063 | -35.8442 | -35.5311
24.0691 | -35.8447 | -35.537
24.0752 | -35.8452 | -35.5438
24.0813 | -35.8457 | -35.5499
24.0874 | -35.8462 | -35.5562
24.0935 | -35.8467 | -35.5623
24.0996 | -35.8471 | -35.5678
24.1057 | -35.8476 | -35.5735
24.1118 | -35.8481 | -35.5797
24.1179 | -35.8485 | -35.5857
24.124 | -35.849 | -35.5914
24.1301 | -35.8494 | -35.5974
24.1362 | -35.8499 | -35.6033
24.1424 | -35.8504 | -35.6087
24.1485 | -35.8509 | -35.6149
24.1546 | -35.8514 | -35.6211
24.1607 | -35.8519 | -35.627
24.1668 | -35.8524 | -35.6329
24.1729 | -35.8528 | -35.6391
24.179 | -35.8532 | -35.6448
24.1851 | -35.8537 | -35.6507
24.1912 | -35.8543 | -35.6569
24.1973 | -35.8548 | -35.6632
24.2034 | -35.8553 | -35.669
24.2095 | -35.8558 | -35.6747
24.2156 | -35.8563 | -35.6809
24.2217 | -35.8567 | -35.6869
24.2278 | -35.8572 | -35.6928
24.2339 | -35.8577 | -35.6985
24.24 | -35.8582 | -35.7044
24.2461 | -35.8587 | -35.7103
24.2522 | -35.8592 | -35.716
24.2583 | -35.8597 | -35.7221
24.2644 | -35.8601 | -35.7278
24.2705 | -35.8606 | -35.7335
24.2766 | -35.8611 | -35.7392
24.2827 | -35.8615 | -35.7455
24.2889 | -35.862 | -35.7515
24.295 | -35.8625 | -35.7578
24.3011 | -35.863 | -35.764
24.3072 | -35.8635 | -35.7704
24.3133 | -35.864 | -35.7762
24.3194 | -35.8644 | -35.7821
24.3255 | -35.8648 | -35.7878
24.3316 | -35.8653 | -35.7937
24.3377 | -35.8658 | -35.7995
24.3438 | -35.8663 | -35.8056
24.3499 | -35.8667 | -35.8112
24.356 | -35.8672 | -35.8175
24.3621 | -35.8676 | -35.824
24.3682 | -35.8681 | -35.8303
24.3743 | -35.8686 | -35.8365
24.3804 | -35.869 | -35.8424
24.3865 | -35.8695 | -35.8486
24.3926 | -35.8699 | -35.8544
24.3987 | -35.8704 | -35.8608
24.4048 | -35.8709 | -35.8668
24.4109 | -35.8713 | -35.873
24.417 | -35.8718 | -35.8787
24.4231 | -35.8723 | -35.8848
24.4293 | -35.8727 | -35.8907
24.4354 | -35.8732 | -35.8964
24.4415 | -35.8737 | -35.9022
24.4476 | -35.8742 | -35.908
24.4537 | -35.8747 | -35.9135
24.4598 | -35.8752 | -35.9197
24.4659 | -35.8757 | -35.9256
24.472 | -35.8762 | -35.9313
24.4781 | -35.8766 | -35.9371
24.4842 | -35.877 | -35.9429
24.4903 | -35.8776 | -35.9489
24.4964 | -35.878 | -35.955
24.5025 | -35.8785 | -35.9611
24.5086 | -35.879 | -35.9668
24.5147 | -35.8795 | -35.9727
24.5208 | -35.8801 | -35.9789
24.5269 | -35.8806 | -35.9845
24.533 | -35.8811 | -35.9903
24.5391 | -35.8816 | -35.9966
24.5452 | -35.8821 | -36.0023
24.5513 | -35.8826 | -36.0078
24.5574 | -35.8831 | -36.0138
24.5635 | -35.8836 | -36.0195
24.5696 | -35.8841 | -36.0252
24.5758 | -35.8847 | -36.0307
24.5819 | -35.8851 | -36.037
24.588 | -35.8856 | -36.0426
24.5941 | -35.8861 | -36.0485
24.6002 | -35.8866 | -36.0544
24.6063 | -35.8871 | -36.0603
24.6124 | -35.8875 | -36.0664
24.6185 | -35.8879 | -36.0722
24.6246 | -35.8884 | -36.0779
24.6307 | -35.8889 | -36.084
24.6368 | -35.8893 | -36.09
24.6429 | -35.8898 | -36.0959
24.649 | -35.8903 | -36.1014
24.6551 | -35.8908 | -36.1072
24.6612 | -35.8913 | -36.1132
24.6673 | -35.8918 | -36.1192
24.6734 | -35.8922 | -36.125
24.6795 | -35.8928 | -36.1312
24.6856 | -35.8933 | -36.1372
24.6917 | -35.8938 | -36.1428
24.6978 | -35.8943 | -36.1489
24.7039 | -35.8948 | -36.1549
24.71 | -35.8953 | -36.1606
24.7162 | -35.8958 | -36.1665
24.7223 | -35.8964 | -36.1722
24.7284 | -35.8968 | -36.1781
24.7345 | -35.8974 | -36.1841
24.7406 | -35.8979 | -36.1906
24.7467 | -35.8984 | -36.1965
24.7528 | -35.8989 | -36.2023
24.7589 | -35.8993 | -36.2081
24.765 | -35.8997 | -36.2136
24.7711 | -35.9002 | -36.2197
24.7772 | -35.9007 | -36.2255
24.7833 | -35.9012 | -36.2314
24.7894 | -35.9016 | -36.2371
24.7955 | -35.9021 | -36.2429
24.8016 | -35.9026 | -36.248
24.8077 | -35.9031 | -36.254
24.8138 | -35.9036 | -36.2601
24.8199 | -35.9041 | -36.2659
24.826 | -35.9046 | -36.2716
24.8321 | -35.9051 | -36.2774
24.8382 | -35.9055 | -36.283
24.8443 | -35.906 | -36.2893
24.8504 | -35.9066 | -36.2952
24.8565 | -35.907 | -36.3011
24.8627 | -35.9075 | -36.3072
24.8688 | -35.9079 | -36.3131
24.8749 | -35.9084 | -36.3191
24.881 | -35.9088 | -36.3248
24.8871 | -35.9093 | -36.3308
24.8932 | -35.9098 | -36.3363
24.8993 | -35.9103 | -36.3419
24.9054 | -35.9107 | -36.3483
24.9115 | -35.9112 | -36.3543
24.9176 | -35.9116 | -36.3602
24.9237 | -35.9121 | -36.3658
24.9298 | -35.9125 | -36.372
24.9359 | -35.913 | -36.3781
24.942 | -35.9135 | -36.3838
24.9481 | -35.914 | -36.3895
24.9542 | -35.9144 | -36.3955
24.9603 | -35.9148 | -36.4014
24.9664 | -35.9153 | -36.4073
24.9725 | -35.9158 | -36.4135
24.9786 | -35.9163 | -36.4189
24.9847 | -35.9168 | -36.4247
24.9908 | -35.9172 | -36.4307
24.9969 | -35.9177 | -36.4365
25.0031 | -35.9182 | -36.4425
25.0092 | -35.9186 | -36.4484
25.0153 | -35.9191 | -36.4539
25.0214 | -35.9195 | -36.4604
25.0275 | -35.92 | -36.4661
25.0336 | -35.9205 | -36.4717
25.0397 | -35.921 | -36.4778
25.0458 | -35.9214 | -36.4835
25.0519 | -35.9219 | -36.4896
25.058 | -35.9224 | -36.4958
25.0641 | -35.9228 | -36.5014
25.0702 | -35.9233 | -36.5074
25.0763 | -35.9238 | -36.5131
25.0824 | -35.9243 | -36.5195
25.0885 | -35.9247 | -36.5252
25.0946 | -35.9252 | -36.531
25.1007 | -35.9257 | -36.537
25.1068 | -35.9261 | -36.5431
25.1129 | -35.9266 | -36.549
25.119 | -35.9271 | -36.5549
25.1251 | -35.9276 | -36.5608
25.1312 | -35.928 | -36.5663
25.1373 | -35.9284 | -36.5724
25.1435 | -35.9289 | -36.5782
25.1496 | -35.9294 | -36.584
25.1557 | -35.9298 | -36.59
25.1618 | -35.9303 | -36.5959
25.1679 | -35.9309 | -36.6017
25.174 | -35.9313 | -36.6078
25.1801 | -35.9319 | -36.6137
25.1862 | -35.9324 | -36.6194
25.1923 | -35.9328 | -36.6252
25.1984 | -35.9334 | -36.6313
25.2045 | -35.9339 | -36.6376
25.2106 | -35.9343 | -36.6434
25.2167 | -35.9348 | -36.6494
25.2228 | -35.9353 | -36.6552
25.2289 | -35.9358 | -36.661
25.235 | -35.9363 | -36.6668
25.2411 | -35.9367 | -36.6727
25.2472 | -35.9372 | -36.6786
25.2533 | -35.9376 | -36.6849
25.2594 | -35.9381 | -36.6903
25.2655 | -35.9386 | -36.6963
25.2716 | -35.939 | -36.7018
25.2777 | -35.9395 | -36.7075
25.2838 | -35.94 | -36.7135
25.29 | -35.9405 | -36.7194
25.2961 | -35.9409 | -36.7254
25.3022 | -35.9415 | -36.7318
25.3083 | -35.9419 | -36.7372
25.3144 | -35.9424 | -36.7434
25.3205 | -35.9428 | -36.749
25.3266 | -35.9433 | -36.7543
25.3327 | -35.9438 | -36.7603
25.3388 | -35.9442 | -36.7658
25.3449 | -35.9447 | -36.7719
25.351 | -35.9452 | -36.7778
25.3571 | -35.9457 | -36.783
25.3632 | -35.9462 | -36.7886
25.3693 | -35.9467 | -36.7947
25.3754 | -35.9472 | -36.8006
25.3815 | -35.9477 | -36.8061
25.3876 | -35.9481 | -36.8118
25.3937 | -35.9486 | -36.8172
25.3998 | -35.949 | -36.8231
25.4059 | -35.9495 | -36.8293
25.412 | -35.9499 | -36.8346
25.4181 | -35.9503 | -36.8406
25.4242 | -35.9508 | -36.8465
25.4304 | -35.9513 | -36.8522
25.4365 | -35.9516 | -36.8577
25.4426 | -35.952 | -36.8635
25.4487 | -35.9525 | -36.869
25.4548 | -35.953 | -36.875
25.4609 | -35.9535 | -36.8806
25.467 | -35.9539 | -36.8857
25.4731 | -35.9544 | -36.8916
25.4792 | -35.9549 | -36.8971
25.4853 | -35.9553 | -36.9032
25.4914 | -35.9557 | -36.9091
25.4975 | -35.9562 | -36.9148
25.5036 | -35.9567 | -36.9209
25.5097 | -35.9572 | -36.9265
25.5158 | -35.9577 | -36.9322
25.5219 | -35.9582 | -36.9383
25.528 | -35.9587 | -36.9442
25.5341 | -35.9591 | -36.95
25.5402 | -35.9595 | -36.9557
25.5463 | -35.96 | -36.9616
25.5524 | -35.9605 | -36.9678
25.5585 | -35.961 | -36.9737
25.5646 | -35.9615 | -36.9794
25.5707 | -35.9619 | -36.9855
25.5769 | -35.9623 | -36.9917
25.583 | -35.9627 | -36.9974
25.5891 | -35.9632 | -37.0032
25.5952 | -35.9637 | -37.0091
25.6013 | -35.9641 | -37.0149
25.6074 | -35.9646 | -37.0202
25.6135 | -35.9651 | -37.0262
25.6196 | -35.9656 | -37.0321
25.6257 | -35.966 | -37.0377
25.6318 | -35.9664 | -37.0435
25.6379 | -35.9669 | -37.0495
25.644 | -35.9673 | -37.0551
25.6501 | -35.9678 | -37.0609
25.6562 | -35.9683 | -37.067
25.6623 | -35.9688 | -37.0727
25.6684 | -35.9693 | -37.0785
25.6745 | -35.9697 | -37.0841
25.6806 | -35.9701 | -37.0897
25.6867 | -35.9705 | -37.0957
25.6928 | -35.9709 | -37.1012
25.6989 | -35.9714 | -37.107
25.705 | -35.9718 | -37.1131
25.7111 | -35.9723 | -37.1192
25.7173 | -35.9727 | -37.125
25.7234 | -35.9732 | -37.131
25.7295 | -35.9737 | -37.1372
25.7356 | -35.9741 | -37.1432
25.7417 | -35.9746 | -37.1488
25.7478 | -35.975 | -37.1541
25.7539 | -35.9754 | -37.1602
25.76 | -35.976 | -37.1661
25.7661 | -35.9764 | -37.1718
25.7722 | -35.9768 | -37.1774
25.7783 | -35.9773 | -37.1832
25.7844 | -35.9779 | -37.1888
25.7905 | -35.9783 | -37.1948
25.7966 | -35.9787 | -37.2006
25.8027 | -35.9791 | -37.2064
25.8088 | -35.9796 | -37.2127
25.8149 | -35.9801 | -37.218
25.821 | -35.9805 | -37.2238
25.8271 | -35.981 | -37.2299
25.8332 | -35.9814 | -37.236
25.8393 | -35.9819 | -37.2415
25.8454 | -35.9824 | -37.2469
25.8515 | -35.9829 | -37.2525
25.8576 | -35.9834 | -37.2587
25.8638 | -35.9838 | -37.2645
25.8699 | -35.9842 | -37.2704
25.876 | -35.9847 | -37.2765
25.8821 | -35.9852 | -37.2825
25.8882 | -35.9856 | -37.2877
25.8943 | -35.9861 | -37.2937
25.9004 | -35.9865 | -37.2996
25.9065 | -35.987 | -37.3056
25.9126 | -35.9874 | -37.311
25.9187 | -35.9879 | -37.3167
25.9248 | -35.9884 | -37.3225
25.9309 | -35.9888 | -37.3284
25.937 | -35.9894 | -37.3341
25.9431 | -35.9899 | -37.3393
25.9492 | -35.9904 | -37.3451
25.9553 | -35.9908 | -37.3511
25.9614 | -35.9913 | -37.3567
25.9675 | -35.9916 | -37.3621
25.9736 | -35.9921 | -37.368
25.9797 | -35.9926 | -37.3739
25.9858 | -35.993 | -37.3796
25.9919 | -35.9936 | -37.386
25.998 | -35.994 | -37.3912
26.0042 | -35.9945 | -37.3968
26.0103 | -35.995 | -37.4021
26.0164 | -35.9954 | -37.408
26.0225 | -35.9958 | -37.414
26.0286 | -35.9963 | -37.4196
26.0347 | -35.9968 | -37.4253
26.0408 | -35.9972 | -37.4306
26.0469 | -35.9976 | -37.4361
26.053 | -35.998 | -37.4415
26.0591 | -35.9984 | -37.4479
26.0652 | -35.9989 | -37.4533
26.0713 | -35.9993 | -37.4594
26.0774 | -35.9998 | -37.465
26.0835 | -36.0003 | -37.4708
26.0896 | -36.0008 | -37.4765
26.0957 | -36.0013 | -37.4826
26.1018 | -36.0016 | -37.4884
26.1079 | -36.0021 | -37.4942
26.114 | -36.0024 | -37.5002
26.1201 | -36.0028 | -37.5055
26.1262 | -36.0033 | -37.5112
26.1323 | -36.0037 | -37.5163
26.1384 | -36.0042 | -37.5219
26.1445 | -36.0046 | -37.5274
26.1507 | -36.005 | -37.5333
26.1568 | -36.0054 | -37.5391
26.1629 | -36.0058 | -37.5448
26.169 | -36.0063 | -37.5508
26.1751 | -36.0066 | -37.5572
26.1812 | -36.0071 | -37.5631
26.1873 | -36.0075 | -37.569
26.1934 | -36.008 | -37.5743
26.1995 | -36.0085 | -37.5801
26.2056 | -36.009 | -37.5859
26.2117 | -36.0094 | -37.5918
26.2178 | -36.0099 | -37.5975
26.2239 | -36.0104 | -37.603
26.23 | -36.0109 | -37.6086
26.2361 | -36.0113 | -37.6143
26.2422 | -36.0119 | -37.6196
26.2483 | -36.0124 | -37.6251
26.2544 | -36.0129 | -37.6306
26.2605 | -36.0133 | -37.6366
26.2666 | -36.0138 | -37.6427
26.2727 | -36.0141 | -37.6482
26.2788 | -36.0145 | -37.653
26.2849 | -36.015 | -37.6589
26.2911 | -36.0154 | -37.6642
26.2972 | -36.0159 | -37.6701
26.3033 | -36.0164 | -37.6758
26.3094 | -36.0167 | -37.682
26.3155 | -36.0171 | -37.6882
26.3216 | -36.0175 | -37.6937
26.3277 | -36.0179 | -37.6988
26.3338 | -36.0184 | -37.7046
26.3399 | -36.0188 | -37.7106
26.346 | -36.0192 | -37.7161
26.3521 | -36.0196 | -37.7219
26.3582 | -36.02 | -37.7278
26.3643 | -36.0204 | -37.7333
26.3704 | -36.0209 | -37.7386
26.3765 | -36.0214 | -37.744
26.3826 | -36.0218 | -37.7503
26.3887 | -36.0223 | -37.756
26.3948 | -36.0228 | -37.7616
26.4009 | -36.0232 | -37.7676
26.407 | -36.0236 | -37.7736
26.4131 | -36.0241 | -37.7796
26.4192 | -36.0244 | -37.7857
26.4253 | -36.0249 | -37.7908
26.4314 | -36.0253 | -37.7969
26.4376 | -36.0257 | -37.8028
26.4437 | -36.0261 | -37.8085
26.4498 | -36.0266 | -37.8137
26.4559 | -36.0271 | -37.8196
26.462 | -36.0275 | -37.8255
26.4681 | -36.0279 | -37.8315
26.4742 | -36.0284 | -37.8375
26.4803 | -36.0289 | -37.843
26.4864 | -36.0293 | -37.8487
26.4925 | -36.0297 | -37.8542
26.4986 | -36.0302 | -37.86
26.5047 | -36.0306 | -37.8661
26.5108 | -36.0311 | -37.8711
26.5169 | -36.0315 | -37.8767
26.523 | -36.0318 | -37.8829
26.5291 | -36.0323 | -37.8888
26.5352 | -36.0327 | -37.8942
26.5413 | -36.0332 | -37.9003
26.5474 | -36.0337 | -37.9059
26.5535 | -36.0342 | -37.9118
26.5596 | -36.0346 | -37.9171
26.5657 | -36.035 | -37.9225
26.5718 | -36.0355 | -37.9284
26.578 | -36.0358 | -37.9345
26.5841 | -36.0362 | -37.9403
26.5902 | -36.0368 | -37.9462
26.5963 | -36.0372 | -37.9517
26.6024 | -36.0376 | -37.9578
26.6085 | -36.0382 | -37.9635
26.6146 | -36.0386 | -37.9692
26.6207 | -36.039 | -37.9754
26.6268 | -36.0395 | -37.9811
26.6329 | -36.0399 | -37.9874
26.639 | -36.0403 | -37.9935
26.6451 | -36.0407 | -37.9986
26.6512 | -36.0411 | -38.004
26.6573 | -36.0415 | -38.0096
26.6634 | -36.0419 | -38.0155
26.6695 | -36.0424 | -38.0207
26.6756 | -36.0428 | -38.0264
26.6817 | -36.0431 | -38.0316
26.6878 | -36.0435 | -38.0374
26.6939 | -36.0439 | -38.0425
26.7 | -36.0444 | -38.0483
26.7061 | -36.0448 | -38.0538
26.7122 | -36.0452 | -38.0593
26.7183 | -36.0456 | -38.0651
26.7245 | -36.0461 | -38.0708
26.7306 | -36.0466 | -38.0763
26.7367 | -36.047 | -38.0809
26.7428 | -36.0474 | -38.0864
26.7489 | -36.0479 | -38.0926
26.755 | -36.0483 | -38.0977
26.7611 | -36.0488 | -38.1034
26.7672 | -36.0493 | -38.11
26.7733 | -36.0498 | -38.1152
26.7794 | -36.0502 | -38.121
26.7855 | -36.0507 | -38.1261
26.7916 | -36.051 | -38.1318
26.7977 | -36.0514 | -38.1376
26.8038 | -36.052 | -38.1434
26.8099 | -36.0524 | -38.1492
26.816 | -36.0528 | -38.1546
26.8221 | -36.0532 | -38.1604
26.8282 | -36.0536 | -38.1657
26.8343 | -36.054 | -38.1719
26.8404 | -36.0545 | -38.1769
26.8465 | -36.0549 | -38.1831
26.8526 | -36.0553 | -38.1881
26.8587 | -36.0557 | -38.1939
26.8649 | -36.0562 | -38.1999
26.871 | -36.0566 | -38.2058
26.8771 | -36.0571 | -38.2113
26.8832 | -36.0576 | -38.2166
26.8893 | -36.058 | -38.2225
26.8954 | -36.0585 | -38.2285
26.9015 | -36.0591 | -38.2339
26.9076 | -36.0594 | -38.2408
26.9137 | -36.0599 | -38.2462
26.9198 | -36.0604 | -38.252
26.9259 | -36.0608 | -38.2574
26.932 | -36.0612 | -38.263
26.9381 | -36.0616 | -38.2693
26.9442 | -36.0621 | -38.2743
26.9503 | -36.0624 | -38.2799
26.9564 | -36.0629 | -38.2857
26.9625 | -36.0633 | -38.2912
26.9686 | -36.0638 | -38.2972
26.9747 | -36.0642 | -38.3026
26.9808 | -36.0646 | -38.308
26.9869 | -36.065 | -38.3134
26.993 | -36.0653 | -38.3188
26.9991 | -36.0659 | -38.3241
27.0052 | -36.0663 | -38.3303
27.0114 | -36.0667 | -38.3361
27.0175 | -36.067 | -38.3422
27.0236 | -36.0674 | -38.3478
27.0297 | -36.0678 | -38.3535
27.0358 | -36.0681 | -38.3588
27.0419 | -36.0685 | -38.3649
27.048 | -36.0689 | -38.3709
27.0541 | -36.0693 | -38.3765
27.0602 | -36.0697 | -38.3822
27.0663 | -36.0701 | -38.3878
27.0724 | -36.0705 | -38.3933
27.0785 | -36.071 | -38.3994
27.0846 | -36.0713 | -38.4042
27.0907 | -36.0717 | -38.4105
27.0968 | -36.0721 | -38.4165
27.1029 | -36.0725 | -38.422
27.109 | -36.0731 | -38.4273
27.1151 | -36.0735 | -38.433
27.1212 | -36.0738 | -38.438
27.1273 | -36.0743 | -38.4438
27.1334 | -36.0747 | -38.4497
27.1395 | -36.0751 | -38.4557
27.1456 | -36.0754 | -38.4611
27.1518 | -36.0759 | -38.4664
27.1579 | -36.0763 | -38.4716
27.164 | -36.0767 | -38.477
27.1701 | -36.0772 | -38.4824
27.1762 | -36.0776 | -38.4887
27.1823 | -36.078 | -38.4942
27.1884 | -36.0785 | -38.5
27.1945 | -36.079 | -38.5062
27.2006 | -36.0795 | -38.5118
27.2067 | -36.0799 | -38.5173
27.2128 | -36.0803 | -38.5234
27.2189 | -36.0807 | -38.5285
27.225 | -36.0812 | -38.534
27.2311 | -36.0815 | -38.5395
27.2372 | -36.082 | -38.5448
27.2433 | -36.0824 | -38.5499
27.2494 | -36.0828 | -38.5556
27.2555 | -36.0831 | -38.5616
27.2616 | -36.0834 | -38.5674
27.2677 | -36.0838 | -38.5728
27.2738 | -36.0843 | -38.5783
27.2799 | -36.0848 | -38.5842
27.286 | -36.0852 | -38.5896
27.2921 | -36.0856 | -38.5949
27.2983 | -36.0859 | -38.6004
27.3044 | -36.0863 | -38.606
27.3105 | -36.0867 | -38.6115
27.3166 | -36.0871 | -38.6171
27.3227 | -36.0876 | -38.6226
27.3288 | -36.0881 | -38.628
27.3349 | -36.0885 | -38.6338
27.341 | -36.0889 | -38.6391
27.3471 | -36.0892 | -38.6441
27.3532 | -36.0896 | -38.65
27.3593 | -36.0901 | -38.656
27.3654 | -36.0906 | -38.6615
27.3715 | -36.091 | -38.667
27.3776 | -36.0913 | -38.6732
27.3837 | -36.0918 | -38.6789
27.3898 | -36.0921 | -38.6846
27.3959 | -36.0926 | -38.6901
27.402 | -36.0929 | -38.6952
27.4081 | -36.0932 | -38.7012
27.4142 | -36.0936 | -38.7072
27.4203 | -36.094 | -38.7129
27.4264 | -36.0945 | -38.7183
27.4325 | -36.0949 | -38.7246
27.4387 | -36.0953 | -38.7311
27.4448 | -36.0957 | -38.7365
27.4509 | -36.0962 | -38.7413
27.457 | -36.0967 | -38.7467
27.4631 | -36.0972 | -38.7527
27.4692 | -36.0975 | -38.7581
27.4753 | -36.0979 | -38.7629
27.4814 | -36.0982 | -38.7688
27.4875 | -36.0987 | -38.7745
27.4936 | -36.0991 | -38.7799
27.4997 | -36.0995 | -38.7855
27.5058 | -36.0999 | -38.7912
27.5119 | -36.1004 | -38.7963
27.518 | -36.1007 | -38.8015
27.5241 | -36.1011 | -38.8069
27.5302 | -36.1015 | -38.8131
27.5363 | -36.1019 | -38.819
27.5424 | -36.1025 | -38.8245
27.5485 | -36.1028 | -38.8303
27.5546 | -36.1031 | -38.8355
27.5607 | -36.1035 | -38.8422
27.5668 | -36.1038 | -38.848
27.5729 | -36.1042 | -38.8538
27.5791 | -36.1046 | -38.8584
27.5852 | -36.1051 | -38.864
27.5913 | -36.1056 | -38.8699
27.5974 | -36.1059 | -38.876
27.6035 | -36.1063 | -38.8814
27.6096 | -36.1067 | -38.8872
27.6157 | -36.1072 | -38.8923
27.6218 | -36.1076 | -38.8978
27.6279 | -36.108 | -38.904
27.634 | -36.1084 | -38.9101
27.6401 | -36.1088 | -38.9156
27.6462 | -36.1092 | -38.9206
27.6523 | -36.1095 | -38.9257
27.6584 | -36.1099 | -38.9309
27.6645 | -36.1103 | -38.9365
27.6706 | -36.1106 | -38.942
27.6767 | -36.111 | -38.9477
27.6828 | -36.1114 | -38.9531
27.6889 | -36.1119 | -38.9591
27.695 | -36.1122 | -38.9648
27.7011 | -36.1126 | -38.9706
27.7072 | -36.1131 | -38.9762
27.7133 | -36.1136 | -38.9816
27.7194 | -36.1139 | -38.9855
27.7256 | -36.1144 | -38.9909
27.7317 | -36.1147 | -38.9966
27.7378 | -36.1151 | -39.0015
27.7439 | -36.1156 | -39.0068
27.75 | -36.1161 | -39.0127
27.7561 | -36.1164 | -39.0186
27.7622 | -36.1168 | -39.0236
27.7683 | -36.1171 | -39.0287
27.7744 | -36.1175 | -39.0344
27.7805 | -36.1179 | -39.04
27.7866 | -36.1183 | -39.0461
27.7927 | -36.1187 | -39.0511
27.7988 | -36.1191 | -39.0565
27.8049 | -36.1195 | -39.0616
27.811 | -36.1199 | -39.0667
27.8171 | -36.1203 | -39.072
27.8232 | -36.1208 | -39.0765
27.8293 | -36.1212 | -39.0814
27.8354 | -36.1216 | -39.0871
27.8415 | -36.122 | -39.0929
27.8476 | -36.1224 | -39.0988
27.8537 | -36.1228 | -39.1044
27.8598 | -36.1232 | -39.11
27.866 | -36.1237 | -39.1157
27.8721 | -36.1242 | -39.121
27.8782 | -36.1245 | -39.1265
27.8843 | -36.125 | -39.1321
27.8904 | -36.1255 | -39.1377
27.8965 | -36.1259 | -39.143
27.9026 | -36.1263 | -39.1477
27.9087 | -36.1267 | -39.1525
27.9148 | -36.1271 | -39.1579
27.9209 | -36.1275 | -39.1635
27.927 | -36.1278 | -39.1692
27.9331 | -36.1282 | -39.1753
27.9392 | -36.1286 | -39.1808
27.9453 | -36.129 | -39.1864
27.9514 | -36.1293 | -39.1914
27.9575 | -36.1297 | -39.1972
27.9636 | -36.1302 | -39.2029
27.9697 | -36.1306 | -39.2097
27.9758 | -36.131 | -39.2151
27.9819 | -36.1314 | -39.2203
27.988 | -36.1318 | -39.2257
27.9941 | -36.1322 | -39.2311
28.0002 | -36.1325 | -39.2382
28.0063 | -36.133 | -39.2435
28.0125 | -36.1333 | -39.2487
28.0186 | -36.1338 | -39.2536
28.0247 | -36.1343 | -39.2589
28.0308 | -36.1347 | -39.2639
28.0369 | -36.1351 | -39.2695
28.043 | -36.1355 | -39.2753
28.0491 | -36.1359 | -39.2814
28.0552 | -36.1363 | -39.2865
28.0613 | -36.1367 | -39.2931
28.0674 | -36.1371 | -39.2987
28.0735 | -36.1374 | -39.3036

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.7452 | -43.8084
21.4908 | -35.7613 | -36.4581
16.4948 | -35.1049 | -28.8908
12.6602 | -34.7176 | -22.4628
9.7171 | -34.4778 | -17.4401
7.4582 | -34.3121 | -13.3684
5.7244 | -34.1916 | -10.0666
4.3936 | -34.1069 | -7.4054
3.3722 | -34.0838 | -5.3026
2.5883 | -34.0684 | -3.6999
1.9866 | -34.0652 | -2.4224
1.5248 | -34.0486 | -1.3876
1.1703 | -34.0519 | -0.71011
1 | -34.047 | -0.37398
0.89822 | -34.0388 | -0.094267
0.68941 | -34.017 | 0.50148
0.52919 | -34.0092 | 0.9767
0.40612 | -33.9953 | 1.6098
0.31171 | -33.9909 | 2.0743
0.23928 | -33.9813 | 2.8186
0.18358 | -34.0001 | 3.2402
0.14091 | -33.9883 | 4.1343
0.10819 | -33.9974 | 5.2999
0.083017 | -34.0189 | 6.776
0.063705 | -34.0637 | 8.6611
0.048923 | -34.1309 | 11.1685
0.037479 | -34.229 | 14.3712
0.028801 | -34.41 | 18.4244
0.022078 | -34.7188 | 23.817
0.017023 | -35.2005 | 30.2387
0.013018 | -36.0017 | 37.9154
0.010014 | -37.1071 | 45.8174

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003666 | 2.1479
0.007331 | -29.744
0.010997 | -36.6782
0.014662 | -41.1051
0.018328 | -44.0422
0.021993 | -45.5739
0.025659 | -45.9837
0.029324 | -47.9634
0.03299 | -48.9891
0.036655 | -49.538
0.040321 | -51.334
0.043986 | -52.7018
0.047652 | -53.7343
0.051317 | -53.9272
0.054983 | -55.1536
0.058648 | -55.1984
0.062314 | -55.6865
0.065979 | -58.2207
0.069645 | -58.1623
0.07331 | -58.1344
0.076976 | -58.1861
0.080641 | -59.7283
0.084307 | -61.3166
0.087973 | -61.5846
0.091638 | -61.1923
0.095304 | -61.8038
0.098969 | -62.6869
0.10264 | -63.5039
0.1063 | -63.7056
0.10997 | -63.6935
0.11363 | -64.2548
0.1173 | -64.6085
0.12096 | -65.407
0.12463 | -64.8228
0.12829 | -63.932
0.13196 | -64.678
0.13562 | -65.6479
0.13929 | -66.4231
0.14295 | -67.0023
0.14662 | -67.3906
0.15029 | -67.5249
0.15395 | -67.8104
0.15762 | -68.0719
0.16128 | -67.6923
0.16495 | -68.2451
0.16861 | -69.231
0.17228 | -68.7642
0.17594 | -69.5508
0.17961 | -69.3733
0.18328 | -68.8524
0.18694 | -69.7548
0.19061 | -69.8726
0.19427 | -69.22
0.19794 | -69.1166
0.2016 | -69.7954
0.20527 | -69.3547
0.20894 | -70.2079
0.2126 | -71.1945
0.21627 | -71.1298
0.21993 | -72.6483
0.2236 | -73.3495
0.22726 | -72.0184
0.23093 | -72.8463
0.23459 | -72.9062
0.23826 | -72.8174
0.24192 | -72.3222
0.24559 | -73.1711
0.24926 | -74.2165
0.25292 | -73.9924
0.25659 | -73.9482
0.26025 | -72.9852
0.26392 | -73.0578
0.26758 | -72.8795
0.27125 | -73.0659
0.27491 | -74.0939
0.27858 | -73.5312
0.28225 | -73.157
0.28591 | -74.1724
0.28958 | -74.5722
0.29324 | -74.6937
0.29691 | -74.9048
0.30057 | -74.5514
0.30424 | -74.1964
0.3079 | -74.0802
0.31157 | -74.3682
0.31523 | -74.7243
0.3189 | -75.3023
0.32257 | -75.9517
0.32623 | -75.6615
0.3299 | -75.6642
0.33356 | -76.4319
0.33723 | -76.4227
0.34089 | -76.5834
0.34456 | -75.9214
0.34823 | -75.252
0.35189 | -75.4242
0.35556 | -76.4747
0.35922 | -77.0467
0.36289 | -77.6457
0.36655 | -77.7327
0.37022 | -77.5237
0.37388 | -76.8085
0.37755 | -77.8068
0.38121 | -77.5906
0.38488 | -78.3916
0.38854 | -78.0092
0.39221 | -77.1239
0.39588 | -77.4947
0.39954 | -78.8631
0.40321 | -79.2756
0.40687 | -79.0155
0.41054 | -77.6711
0.4142 | -78.1079
0.41787 | -78.7771
0.42153 | -79.015
0.4252 | -78.2576
0.42887 | -78.5333
0.43253 | -78.7955
0.4362 | -79.0125
0.43986 | -79.3679
0.44353 | -78.7708
0.44719 | -78.0769
0.45086 | -79.1037
0.45453 | -79.8692
0.45819 | -78.6988
0.46186 | -78.5101
0.46552 | -80.0606
0.46919 | -80.0905
0.47285 | -79.8301
0.47652 | -78.7311
0.48018 | -78.3142
0.48385 | -79.1398
0.48751 | -79.9277
0.49118 | -80.1166
0.49484 | -79.805
0.49851 | -79.6169
0.50218 | -79.2361
0.50584 | -79.9171
0.50951 | -80.6452
0.51317 | -81.6651
0.51684 | -81.4109
0.5205 | -79.4987
0.52417 | -80.2531
0.52784 | -80.7392
0.5315 | -80.6174
0.53517 | -80.459
0.53883 | -80.6192
0.5425 | -81.0274
0.54616 | -81.2279
0.54983 | -81.0096
0.55349 | -81.0868
0.55716 | -81.2652
0.56083 | -81.2821
0.56449 | -81.5106
0.56816 | -81.5368
0.57182 | -81.3967
0.57549 | -81.6093
0.57915 | -81.9689
0.58282 | -82.0263
0.58648 | -81.8551
0.59015 | -81.6676
0.59381 | -81.785
0.59748 | -82.0234
0.60115 | -82.0036
0.60481 | -82.0456
0.60848 | -81.8459
0.61214 | -81.2697
0.61581 | -81.1934
0.61947 | -81.658
0.62314 | -82.0402
0.6268 | -82.4034
0.63047 | -82.7267
0.63414 | -82.7118
0.6378 | -82.395
0.64147 | -81.9502
0.64513 | -81.8246
0.6488 | -82.1585
0.65246 | -82.7892
0.65613 | -83.4143
0.65979 | -83.4734
0.66346 | -83.0304
0.66712 | -82.648
0.67079 | -82.4955
0.67446 | -82.2335
0.67812 | -82.1095
0.68179 | -82.5264
0.68545 | -83.0402
0.68912 | -82.9787
0.69278 | -82.6088
0.69645 | -82.4249
0.70012 | -82.768
0.70378 | -83.1697
0.70745 | -83.1808
0.71111 | -83.0106
0.71478 | -83.0567
0.71844 | -83.5833
0.72211 | -84.3558
0.72577 | -84.2071
0.72944 | -83.3085
0.7331 | -83.0252
0.73677 | -83.029
0.74043 | -83.1943
0.7441 | -83.7743
0.74777 | -84.2831
0.75143 | -84.4477
0.7551 | -84.279
0.75876 | -84.3465
0.76243 | -84.1881
0.76609 | -83.8102
0.76976 | -83.8532
0.77343 | -83.7855
0.77709 | -83.429
0.78076 | -83.05
0.78442 | -83.2346
0.78809 | -83.8482
0.79175 | -84.3542
0.79542 | -84.8187
0.79908 | -85.5292
0.80275 | -86.1672
0.80641 | -85.8315
0.81008 | -85.3565
0.81375 | -85.46
0.81741 | -85.7185
0.82108 | -85.446
0.82474 | -85.208
0.82841 | -85.4137
0.83207 | -85.8289
0.83574 | -86.0594
0.8394 | -85.6699
0.84307 | -85.0707
0.84674 | -84.7703
0.8504 | -84.7067
0.85407 | -84.6741
0.85773 | -85.0406
0.8614 | -85.7036
0.86506 | -86.0549
0.86873 | -85.8149
0.87239 | -85.6968
0.87606 | -86.002
0.87972 | -86.2156
0.88339 | -86.2276
0.88706 | -86.0558
0.89072 | -85.7937
0.89439 | -85.5634
0.89805 | -85.7367
0.90172 | -86.2802
0.90538 | -86.5033
0.90905 | -86.3568
0.91272 | -86.5511
0.91638 | -86.8857
0.92005 | -86.9697
0.92371 | -86.7698
0.92738 | -86.3699
0.93104 | -86.2375
0.93471 | -86.3862
0.93837 | -86.334
0.94204 | -86.457
0.9457 | -86.7125
0.94937 | -86.73
0.95304 | -86.9813
0.9567 | -87.5443
0.96037 | -87.6785
0.96403 | -87.3411
0.9677 | -87.0625
0.97136 | -87.0694
0.97503 | -87.1511
0.97869 | -86.9817
0.98236 | -86.8883
0.98603 | -86.844
0.98969 | -86.0707
0.99336 | -85.5435
0.99702 | -85.9639
1.0007 | -86.7062
1.0044 | -87.3946
1.008 | -87.4189
1.0117 | -86.8199
1.0153 | -86.8189
1.019 | -87.1038
1.0227 | -87.0796
1.0263 | -87.1953
1.03 | -87.5314
1.0337 | -87.6786
1.0373 | -87.9191
1.041 | -87.9074
1.0447 | -87.9803
1.0483 | -88.1877
1.052 | -88.4507
1.0557 | -88.6667
1.0593 | -88.6745
1.063 | -88.407
1.0667 | -87.9825
1.0703 | -87.5837
1.074 | -87.414
1.0777 | -87.4647
1.0813 | -87.5049
1.085 | -87.4664
1.0887 | -87.4625
1.0923 | -87.4237
1.096 | -87.2964
1.0997 | -87.2309
1.1033 | -87.2427
1.107 | -87.2535
1.1107 | -87.327
1.1143 | -87.4971
1.118 | -87.7642
1.1216 | -88.107
1.1253 | -88.4245
1.129 | -88.5504
1.1326 | -88.3965
1.1363 | -88.1313
1.14 | -88.0041
1.1436 | -88.107
1.1473 | -88.4055
1.151 | -88.7478
1.1546 | -88.9185
1.1583 | -88.8822
1.162 | -88.9145
1.1656 | -89.1876
1.1693 | -89.3902
1.173 | -89.1921
1.1766 | -88.8602
1.1803 | -88.7318
1.184 | -88.7921
1.1876 | -88.8521
1.1913 | -88.7176
1.195 | -88.3996
1.1986 | -88.128
1.2023 | -87.9959
1.206 | -87.8584
1.2096 | -87.7247
1.2133 | -87.8314
1.217 | -88.2788
1.2206 | -88.8678
1.2243 | -89.2119
1.228 | -89.1922
1.2316 | -89.0101
1.2353 | -88.6904
1.2389 | -88.3578
1.2426 | -88.2364
1.2463 | -88.3331
1.2499 | -88.5493
1.2536 | -88.7975
1.2573 | -89.054
1.2609 | -89.3232
1.2646 | -89.4575
1.2683 | -89.3252
1.2719 | -89.1247
1.2756 | -89.1119
1.2793 | -89.1871
1.2829 | -89.1709
1.2866 | -89.2082
1.2903 | -89.4262
1.2939 | -89.5978
1.2976 | -89.5085
1.3013 | -89.2714
1.3049 | -89.1413
1.3086 | -89.2373
1.3123 | -89.412
1.3159 | -89.4801
1.3196 | -89.4454
1.3233 | -89.3452
1.3269 | -89.2354
1.3306 | -89.2241
1.3342 | -89.3546
1.3379 | -89.5172
1.3416 | -89.5879
1.3452 | -89.6153
1.3489 | -89.7767
1.3526 | -90.0383
1.3562 | -90.1004
1.3599 | -89.8211
1.3636 | -89.4206
1.3672 | -89.2152
1.3709 | -89.2517
1.3746 | -89.2731
1.3782 | -89.159
1.3819 | -89.1591
1.3856 | -89.4803
1.3892 | -89.8758
1.3929 | -89.9563
1.3966 | -89.6893
1.4002 | -89.3632
1.4039 | -89.3584
1.4076 | -89.7037
1.4112 | -90.0155
1.4149 | -90.0886
1.4186 | -90.1022
1.4222 | -90.2026
1.4259 | -90.371
1.4296 | -90.4446
1.4332 | -90.258
1.4369 | -89.9682
1.4405 | -89.9202
1.4442 | -90.1738
1.4479 | -90.4361
1.4515 | -90.4211
1.4552 | -90.2456
1.4589 | -90.1922
1.4625 | -90.277
1.4662 | -90.3867
1.4699 | -90.5156
1.4735 | -90.511
1.4772 | -90.2624
1.4809 | -90.0288
1.4845 | -90.1042
1.4882 | -90.5385
1.4919 | -91.0765
1.4955 | -91.3361
1.4992 | -91.246
1.5029 | -91.0517
1.5065 | -91.0137
1.5102 | -91.1526
1.5139 | -91.2665
1.5175 | -91.2075
1.5212 | -91.0282
1.5249 | -90.721
1.5285 | -90.2856
1.5322 | -89.8977
1.5359 | -89.7557
1.5395 | -89.9665
1.5432 | -90.407
1.5469 | -90.6912
1.5505 | -90.6712
1.5542 | -90.6109
1.5578 | -90.6965
1.5615 | -90.8603
1.5652 | -91.0259
1.5688 | -91.2317
1.5725 | -91.3635
1.5762 | -91.3053
1.5798 | -91.2404
1.5835 | -91.3809
1.5872 | -91.5491
1.5908 | -91.7809
1.5945 | -92.0113
1.5982 | -92.1637
1.6018 | -92.1948
1.6055 | -92.0978
1.6092 | -91.8917
1.6128 | -91.6091
1.6165 | -91.3333
1.6202 | -91.1888
1.6238 | -91.2122
1.6275 | -91.2945
1.6312 | -91.2948
1.6348 | -91.2201
1.6385 | -91.1889
1.6422 | -91.2511
1.6458 | -91.3361
1.6495 | -91.3545
1.6531 | -91.3135
1.6568 | -91.3307
1.6605 | -91.4747
1.6641 | -91.6827
1.6678 | -91.8314
1.6715 | -91.8191
1.6751 | -91.6873
1.6788 | -91.5777
1.6825 | -91.5431
1.6861 | -91.5167
1.6898 | -91.4373
1.6935 | -91.3586
1.6971 | -91.3865
1.7008 | -91.5359
1.7045 | -91.7404
1.7081 | -91.9467
1.7118 | -92.1623
1.7155 | -92.4181
1.7191 | -92.6821
1.7228 | -92.83
1.7265 | -92.761
1.7301 | -92.4768
1.7338 | -92.0991
1.7375 | -91.7803
1.7411 | -91.6236
1.7448 | -91.6329
1.7485 | -91.7486
1.7521 | -91.9207
1.7558 | -92.1354
1.7594 | -92.3779
1.7631 | -92.6232
1.7668 | -92.7959
1.7704 | -92.8044
1.7741 | -92.6846
1.7778 | -92.6136
1.7814 | -92.6462
1.7851 | -92.6761
1.7888 | -92.6072
1.7924 | -92.4353
1.7961 | -92.2245
1.7998 | -92.0241
1.8034 | -91.8885
1.8071 | -91.8417
1.8108 | -91.8264
1.8144 | -91.7265
1.8181 | -91.5166
1.8218 | -91.3071
1.8254 | -91.2656
1.8291 | -91.4787
1.8328 | -91.9009
1.8364 | -92.3429
1.8401 | -92.6495
1.8438 | -92.8229
1.8474 | -92.9311
1.8511 | -92.9665
1.8548 | -92.8924
1.8584 | -92.7244
1.8621 | -92.549
1.8658 | -92.4793
1.8694 | -92.5742
1.8731 | -92.7637
1.8767 | -92.9198
1.8804 | -92.9904
1.8841 | -93.0315
1.8877 | -93.085
1.8914 | -93.1329
1.8951 | -93.1755
1.8987 | -93.1944
1.9024 | -93.1251
1.9061 | -92.9341
1.9097 | -92.7235
1.9134 | -92.6334
1.9171 | -92.674
1.9207 | -92.7381
1.9244 | -92.7079
1.9281 | -92.6189
1.9317 | -92.6026
1.9354 | -92.6919
1.9391 | -92.8087
1.9427 | -92.8592
1.9464 | -92.8739
1.9501 | -92.907
1.9537 | -92.936
1.9574 | -92.9162
1.9611 | -92.8734
1.9647 | -92.8775
1.9684 | -92.956
1.9721 | -93.0474
1.9757 | -93.1233
1.9794 | -93.1983
1.983 | -93.2679
1.9867 | -93.3059
1.9904 | -93.3426
1.994 | -93.3969
1.9977 | -93.4114
2.0014 | -93.3382
2.005 | -93.2143
2.0087 | -93.0898
2.0124 | -92.9974
2.016 | -92.9711
2.0197 | -93.0019
2.0234 | -93.0489
2.027 | -93.1066
2.0307 | -93.1831
2.0344 | -93.2715
2.038 | -93.3003
2.0417 | -93.2262
2.0454 | -93.1045
2.049 | -93.0223
2.0527 | -93.0673
2.0564 | -93.2252
2.06 | -93.4098
2.0637 | -93.5305
2.0674 | -93.5493
2.071 | -93.5396
2.0747 | -93.5245
2.0784 | -93.4646
2.082 | -93.3996
2.0857 | -93.4119
2.0893 | -93.5388
2.093 | -93.707
2.0967 | -93.8136
2.1003 | -93.7976
2.104 | -93.6524
2.1077 | -93.4393
2.1113 | -93.2796
2.115 | -93.2948
2.1187 | -93.3091
2.1223 | -93.3753
2.126 | -93.468
2.1297 | -93.5575
2.1333 | -93.6242
2.137 | -93.6372
2.1407 | -93.5808
2.1443 | -93.4704
2.148 | -93.3409
2.1517 | -93.2463
2.1553 | -93.2098
2.159 | -93.2009
2.1627 | -93.1838
2.1663 | -93.1259
2.17 | -93.0327
2.1737 | -92.9394
2.1773 | -92.9081
2.181 | -92.9752
2.1847 | -93.113
2.1883 | -93.2615
2.192 | -93.3628
2.1956 | -93.4045
2.1993 | -93.4201
2.203 | -93.4422
2.2066 | -93.4987
2.2103 | -93.5923
2.214 | -93.7333
2.2176 | -93.8951
2.2213 | -94.0471
2.225 | -94.1567
2.2286 | -94.1832
2.2323 | -94.1309
2.236 | -94.0397
2.2396 | -93.9624
2.2433 | -93.9393
2.247 | -93.9665
2.2506 | -94.0113
2.2543 | -94.0517
2.258 | -94.0852
2.2616 | -94.0773
2.2653 | -94.0243
2.269 | -93.951
2.2726 | -93.8852
2.2763 | -93.8659
2.28 | -93.8884
2.2836 | -93.9378
2.2873 | -94.0209
2.291 | -94.145
2.2946 | -94.2629
2.2983 | -94.2793
2.3019 | -94.1557
2.3056 | -93.93
2.3093 | -93.7167
2.3129 | -93.617
2.3166 | -93.6513
2.3203 | -93.8144
2.3239 | -94.0585
2.3276 | -94.297
2.3313 | -94.4324
2.3349 | -94.4248
2.3386 | -94.2983
2.3423 | -94.1102
2.3459 | -93.9413
2.3496 | -93.8528
2.3533 | -93.8639
2.3569 | -93.9618
2.3606 | -94.1131
2.3643 | -94.3134
2.3679 | -94.5407
2.3716 | -94.7426
2.3753 | -94.8901
2.3789 | -94.9778
2.3826 | -95.0323
2.3863 | -95.0709
2.3899 | -95.097
2.3936 | -95.0723
2.3973 | -94.956
2.4009 | -94.7584
2.4046 | -94.5568
2.4082 | -94.4582
2.4119 | -94.5119
2.4156 | -94.6679
2.4192 | -94.8652
2.4229 | -95.0357
2.4266 | -95.1432
2.4302 | -95.1574
2.4339 | -95.0763
2.4376 | -94.9196
2.4412 | -94.73
2.4449 | -94.5525
2.4486 | -94.4442
2.4522 | -94.44
2.4559 | -94.5117
2.4596 | -94.6
2.4632 | -94.6768
2.4669 | -94.7519
2.4706 | -94.8739
2.4742 | -95.0418
2.4779 | -95.2123
2.4816 | -95.3331
2.4852 | -95.401
2.4889 | -95.3977
2.4926 | -95.3021
2.4962 | -95.1585
2.4999 | -95.0223
2.5036 | -94.9462
2.5072 | -94.9505
2.5109 | -94.9849
2.5145 | -94.9836
2.5182 | -94.8969
2.5219 | -94.7193
2.5255 | -94.4921
2.5292 | -94.2801
2.5329 | -94.1425
2.5365 | -94.1217
2.5402 | -94.2498
2.5439 | -94.5193
2.5475 | -94.8725
2.5512 | -95.1922
2.5549 | -95.365
2.5585 | -95.35
2.5622 | -95.179
2.5659 | -94.9335
2.5695 | -94.7026
2.5732 | -94.5384
2.5769 | -94.4569
2.5805 | -94.4283
2.5842 | -94.4334
2.5879 | -94.4859
2.5915 | -94.6156
2.5952 | -94.8273
2.5989 | -95.0961
2.6025 | -95.3846
2.6062 | -95.6431
2.6099 | -95.8205
2.6135 | -95.8995
2.6172 | -95.9002
2.6208 | -95.8627
2.6245 | -95.8067
2.6282 | -95.7061
2.6318 | -95.5546
2.6355 | -95.3833
2.6392 | -95.2085
2.6428 | -95.0972
2.6465 | -95.009
2.6502 | -94.9613
2.6538 | -94.9551
2.6575 | -94.9552
2.6612 | -94.9321
2.6648 | -94.8863
2.6685 | -94.8424
2.6722 | -94.8363
2.6758 | -94.8716
2.6795 | -94.9377
2.6832 | -95.025
2.6868 | -95.1172
2.6905 | -95.185
2.6942 | -95.2086
2.6978 | -95.1946
2.7015 | -95.1441
2.7052 | -95.0631
2.7088 | -94.9622
2.7125 | -94.8582
2.7162 | -94.771
2.7198 | -94.7312
2.7235 | -94.7491
2.7271 | -94.8251
2.7308 | -94.9722
2.7345 | -95.1955
2.7381 | -95.4738
2.7418 | -95.7744
2.7455 | -96.0395
2.7491 | -96.2327
2.7528 | -96.3375
2.7565 | -96.356
2.7601 | -96.2954
2.7638 | -96.1831
2.7675 | -96.0522
2.7711 | -95.9503
2.7748 | -95.8924
2.7785 | -95.8972
2.7821 | -95.9473
2.7858 | -96.0114
2.7895 | -96.0521
2.7931 | -96.039
2.7968 | -95.9572
2.8005 | -95.8165
2.8041 | -95.6353
2.8078 | -95.4485
2.8115 | -95.2951
2.8151 | -95.1845
2.8188 | -95.1286
2.8225 | -95.1417
2.8261 | -95.2008
2.8298 | -95.3074
2.8334 | -95.4451
2.8371 | -95.5852
2.8408 | -95.7136
2.8444 | -95.8157
2.8481 | -95.893
2.8518 | -95.9454
2.8554 | -95.9755
2.8591 | -95.9624
2.8628 | -95.9021
2.8664 | -95.8105
2.8701 | -95.7393
2.8738 | -95.7357
2.8774 | -95.8032
2.8811 | -95.9262
2.8848 | -96.0594
2.8884 | -96.1788
2.8921 | -96.2793
2.8958 | -96.3481
2.8994 | -96.3961
2.9031 | -96.4167
2.9068 | -96.4253
2.9104 | -96.4502
2.9141 | -96.4799
2.9178 | -96.502
2.9214 | -96.4913
2.9251 | -96.4103
2.9288 | -96.2732
2.9324 | -96.1192
2.9361 | -95.9949
2.9397 | -95.9429
2.9434 | -95.9847
2.9471 | -96.0913
2.9507 | -96.2236
2.9544 | -96.3414
2.9581 | -96.4219
2.9617 | -96.4848
2.9654 | -96.5313
2.9691 | -96.586
2.9727 | -96.6476
2.9764 | -96.693
2.9801 | -96.7115
2.9837 | -96.704
2.9874 | -96.6771
2.9911 | -96.6312
2.9947 | -96.5734
2.9984 | -96.4945
3.0021 | -96.4092
3.0057 | -96.3444
3.0094 | -96.322
3.0131 | -96.3535
3.0167 | -96.4185
3.0204 | -96.4972
3.0241 | -96.5775
3.0277 | -96.6468
3.0314 | -96.688
3.0351 | -96.6918
3.0387 | -96.6577
3.0424 | -96.5993
3.046 | -96.5453
3.0497 | -96.5146
3.0534 | -96.5401
3.057 | -96.6241
3.0607 | -96.7295
3.0644 | -96.8047
3.068 | -96.827
3.0717 | -96.8014
3.0754 | -96.7553
3.079 | -96.7197
3.0827 | -96.6945
3.0864 | -96.6632
3.09 | -96.6323
3.0937 | -96.6197
3.0974 | -96.6512
3.101 | -96.7281
3.1047 | -96.8416
3.1084 | -96.9644
3.112 | -97.0484
3.1157 | -97.0558
3.1194 | -96.9787
3.123 | -96.848
3.1267 | -96.6927
3.1304 | -96.5608
3.134 | -96.4752
3.1377 | -96.4447
3.1414 | -96.4763
3.145 | -96.569
3.1487 | -96.7228
3.1523 | -96.919
3.156 | -97.1243
3.1597 | -97.2793
3.1633 | -97.3341
3.167 | -97.2934
3.1707 | -97.1458
3.1743 | -97.0491
3.178 | -96.9703
3.1817 | -96.9154
3.1853 | -96.8975
3.189 | -96.9006
3.1927 | -96.906
3.1963 | -96.8935
3.2 | -96.8367
3.2037 | -96.7354
3.2073 | -96.6265
3.211 | -96.5308
3.2147 | -96.4793
3.2183 | -96.4852
3.222 | -96.5553
3.2257 | -96.682
3.2293 | -96.8372
3.233 | -96.989
3.2367 | -97.1309
3.2403 | -97.2352
3.244 | -97.2871
3.2477 | -97.2909
3.2513 | -97.2629
3.255 | -97.2212
3.2586 | -97.1827
3.2623 | -97.1438
3.266 | -97.1201
3.2696 | -97.0886
3.2733 | -97.0296
3.277 | -96.9529
3.2806 | -96.8705
3.2843 | -96.7997
3.288 | -96.7407
3.2916 | -96.6854
3.2953 | -96.6506
3.299 | -96.633
3.3026 | -96.6322
3.3063 | -96.6531
3.31 | -96.7144
3.3136 | -96.8027
3.3173 | -96.907
3.321 | -97.019
3.3246 | -97.1181
3.3283 | -97.2043
3.332 | -97.2608
3.3356 | -97.2788
3.3393 | -97.2643
3.343 | -97.2324
3.3466 | -97.2045
3.3503 | -97.183
3.354 | -97.1807
3.3576 | -97.2099
3.3613 | -97.2601
3.3649 | -97.3137
3.3686 | -97.3664
3.3723 | -97.4067
3.3759 | -97.426
3.3796 | -97.43
3.3833 | -97.4067
3.3869 | -97.3688
3.3906 | -97.338
3.3943 | -97.3176
3.3979 | -97.3074
3.4016 | -97.3001
3.4053 | -97.2836
3.4089 | -97.2476
3.4126 | -97.1739
3.4163 | -97.0781
3.4199 | -96.9716
3.4236 | -96.8863
3.4273 | -96.8381
3.4309 | -96.8361
3.4346 | -96.8822
3.4383 | -96.9548
3.4419 | -97.0251
3.4456 | -97.0608
3.4493 | -97.054
3.4529 | -97.0011
3.4566 | -96.9266
3.4603 | -96.8683
3.4639 | -96.8486
3.4676 | -96.8671
3.4712 | -96.9176
3.4749 | -96.9933
3.4786 | -97.0673
3.4822 | -97.1236
3.4859 | -97.1526
3.4896 | -97.1539
3.4932 | -97.1473
3.4969 | -97.1406
3.5006 | -97.1457
3.5042 | -97.1674
3.5079 | -97.2067
3.5116 | -97.244
3.5152 | -97.282
3.5189 | -97.3263
3.5226 | -97.386
3.5262 | -97.4531
3.5299 | -97.5065
3.5336 | -97.5383
3.5372 | -97.5374
3.5409 | -97.5143
3.5446 | -97.4592
3.5482 | -97.3789
3.5519 | -97.2944
3.5556 | -97.2232
3.5592 | -97.183
3.5629 | -97.1823
3.5666 | -97.2214
3.5702 | -97.2751
3.5739 | -97.3319
3.5775 | -97.3788
3.5812 | -97.4041
3.5849 | -97.4298
3.5885 | -97.4671
3.5922 | -97.5221
3.5959 | -97.5694
3.5995 | -97.5981
3.6032 | -97.5978
3.6069 | -97.5435
3.6105 | -97.4389
3.6142 | -97.2901
3.6179 | -97.1306
3.6215 | -96.9834
3.6252 | -96.8818
3.6289 | -96.8441
3.6325 | -96.8799
3.6362 | -96.994
3.6399 | -97.1569
3.6435 | -97.3408
3.6472 | -97.531
3.6509 | -97.7082
3.6545 | -97.8499
3.6582 | -97.9247
3.6619 | -97.9187
3.6655 | -97.8264
3.6692 | -97.6767
3.6729 | -97.4912
3.6765 | -97.3164
3.6802 | -97.1752
3.6838 | -97.094
3.6875 | -97.0875
3.6912 | -97.1568
3.6948 | -97.2992
3.6985 | -97.4677
3.7022 | -97.5962
3.7058 | -97.6901
3.7095 | -97.7238
3.7132 | -97.7052
3.7168 | -97.6447
3.7205 | -97.5685
3.7242 | -97.4967
3.7278 | -97.4427
3.7315 | -97.4204
3.7352 | -97.4392
3.7388 | -97.4828
3.7425 | -97.538
3.7462 | -97.5954
3.7498 | -97.6436
3.7535 | -97.6861
3.7572 | -97.7279
3.7608 | -97.7813
3.7645 | -97.8542
3.7682 | -97.9499
3.7718 | -98.0589
3.7755 | -98.1637
3.7792 | -98.2644
3.7828 | -98.3594
3.7865 | -98.4412
3.7901 | -98.5122
3.7938 | -98.5665
3.7975 | -98.6019
3.8011 | -98.6347
3.8048 | -98.6644
3.8085 | -98.6943
3.8121 | -98.7257
3.8158 | -98.7356
3.8195 | -98.735
3.8231 | -98.7231
3.8268 | -98.6948
3.8305 | -98.6505
3.8341 | -98.5894
3.8378 | -98.521
3.8415 | -98.4436
3.8451 | -98.3597
3.8488 | -98.2739
3.8525 | -98.1958
3.8561 | -98.1306
3.8598 | -98.0752
3.8635 | -98.0383
3.8671 | -98.0246
3.8708 | -98.04
3.8745 | -98.0854
3.8781 | -98.1719
3.8818 | -98.2813
3.8855 | -98.4074
3.8891 | -98.5338
3.8928 | -98.6564
3.8964 | -98.7737
3.9001 | -98.8619
3.9038 | -98.9241
3.9074 | -98.9565
3.9111 | -98.949
3.9148 | -98.91
3.9184 | -98.8469
3.9221 | -98.7589
3.9258 | -98.6534
3.9294 | -98.5398
3.9331 | -98.4363
3.9368 | -98.3599
3.9404 | -98.3199
3.9441 | -98.3286
3.9478 | -98.378
3.9514 | -98.4596
3.9551 | -98.5568
3.9588 | -98.6378
3.9624 | -98.6862
3.9661 | -98.684
3.9698 | -98.6436
3.9734 | -98.5773
3.9771 | -98.5001
3.9808 | -98.4446
3.9844 | -98.4364
3.9881 | -98.4765
3.9918 | -98.5568
3.9954 | -98.6546
3.9991 | -98.7517
4.0027 | -98.8224
4.0064 | -98.8527
4.0101 | -98.8459
4.0137 | -98.8133
4.0174 | -98.7543
4.0211 | -98.6922
4.0247 | -98.6332
4.0284 | -98.5855
4.0321 | -98.5539
4.0357 | -98.5496
4.0394 | -98.5662
4.0431 | -98.5979
4.0467 | -98.6395
4.0504 | -98.6874
4.0541 | -98.7386
4.0577 | -98.7896
4.0614 | -98.8381
4.0651 | -98.8843
4.0687 | -98.9108
4.0724 | -98.9116
4.0761 | -98.8798
4.0797 | -98.8166
4.0834 | -98.7263
4.0871 | -98.6301
4.0907 | -98.5453
4.0944 | -98.4733
4.0981 | -98.423
4.1017 | -98.3904
4.1054 | -98.3804
4.109 | -98.3856
4.1127 | -98.4118
4.1164 | -98.4577
4.12 | -98.5042
4.1237 | -98.5532
4.1274 | -98.5994
4.131 | -98.6421
4.1347 | -98.677
4.1384 | -98.7111
4.142 | -98.7543
4.1457 | -98.815
4.1494 | -98.8871
4.153 | -98.9587
4.1567 | -99.026
4.1604 | -99.073
4.164 | -99.0985
4.1677 | -99.0916
4.1714 | -99.0647
4.175 | -99.0337
4.1787 | -99.0027
4.1824 | -98.9885
4.186 | -98.9783
4.1897 | -98.9691
4.1934 | -98.9649
4.197 | -98.9687
4.2007 | -98.9673
4.2044 | -98.9792
4.208 | -98.9935
4.2117 | -99.0119
4.2153 | -99.0268
4.219 | -99.0332
4.2227 | -99.0319
4.2263 | -99.0146
4.23 | -98.9703
4.2337 | -98.9445
4.2373 | -98.9081
4.241 | -98.8684
4.2447 | -98.8329
4.2483 | -98.7963
4.252 | -98.768
4.2557 | -98.7528
4.2593 | -98.745
4.263 | -98.7469
4.2667 | -98.7682
4.2703 | -98.8001
4.274 | -98.8504
4.2777 | -98.9185
4.2813 | -99.0081
4.285 | -99.1113
4.2887 | -99.2058
4.2923 | -99.285
4.296 | -99.3344
4.2997 | -99.3364
4.3033 | -99.2979
4.307 | -99.2303
4.3107 | -99.1436
4.3143 | -99.0489
4.318 | -98.9596
4.3216 | -98.8781
4.3253 | -98.8152
4.329 | -98.7893
4.3326 | -98.7932
4.3363 | -98.8369
4.34 | -98.9059
4.3436 | -98.9807
4.3473 | -99.0523
4.351 | -99.1104
4.3546 | -99.1513
4.3583 | -99.1891
4.362 | -99.2168
4.3656 | -99.2496
4.3693 | -99.2836
4.373 | -99.3151
4.3766 | -99.3375
4.3803 | -99.3453
4.384 | -99.3383
4.3876 | -99.3226
4.3913 | -99.3119
4.395 | -99.313
4.3986 | -99.3131
4.4023 | -99.3304
4.406 | -99.3594
4.4096 | -99.401
4.4133 | -99.4674
4.417 | -99.5347
4.4206 | -99.5962
4.4243 | -99.6572
4.428 | -99.7045
4.4316 | -99.7413
4.4353 | -99.7579
4.4389 | -99.763
4.4426 | -99.7493
4.4463 | -99.7255
4.4499 | -99.6887
4.4536 | -99.6474
4.4573 | -99.6043
4.4609 | -99.5594
4.4646 | -99.51
4.4683 | -99.4685
4.4719 | -99.4268
4.4756 | -99.3892
4.4793 | -99.3661
4.4829 | -99.3536
4.4866 | -99.3678
4.4903 | -99.4035
4.4939 | -99.4468
4.4976 | -99.5045
4.5013 | -99.5593
4.5049 | -99.5972
4.5086 | -99.6178
4.5123 | -99.6191
4.5159 | -99.5955
4.5196 | -99.5655
4.5233 | -99.5333
4.5269 | -99.5115
4.5306 | -99.5097
4.5343 | -99.5285
4.5379 | -99.5606
4.5416 | -99.5945
4.5452 | -99.6246
4.5489 | -99.6421
4.5526 | -99.6417
4.5562 | -99.6202
4.5599 | -99.5874
4.5636 | -99.5546
4.5672 | -99.5156
4.5709 | -99.4788
4.5746 | -99.4519
4.5782 | -99.436
4.5819 | -99.4384
4.5856 | -99.4683
4.5892 | -99.516
4.5929 | -99.5847
4.5966 | -99.6573
4.6002 | -99.7336
4.6039 | -99.8104
4.6076 | -99.874
4.6112 | -99.9169
4.6149 | -99.9352
4.6186 | -99.9385
4.6222 | -99.9325
4.6259 | -99.91
4.6296 | -99.887
4.6332 | -99.8655
4.6369 | -99.8593
4.6406 | -99.8695
4.6442 | -99.8827
4.6479 | -99.8948
4.6515 | -99.901
4.6552 | -99.8882
4.6589 | -99.8631
4.6625 | -99.8199
4.6662 | -99.7658
4.6699 | -99.7101
4.6735 | -99.6612
4.6772 | -99.63
4.6809 | -99.6312
4.6845 | -99.6625
4.6882 | -99.7203
4.6919 | -99.8007
4.6955 | -99.9109
4.6992 | -100.0237
4.7029 | -100.1221
4.7065 | -100.1902
4.7102 | -100.2177
4.7139 | -100.2158
4.7175 | -100.1841
4.7212 | -100.1353
4.7249 | -100.082
4.7285 | -100.0382
4.7322 | -100.0149
4.7359 | -100.0042
4.7395 | -99.9981
4.7432 | -100.0056
4.7469 | -100.0179
4.7505 | -100.0223
4.7542 | -100.0224
4.7578 | -100.0059
4.7615 | -100.0018
4.7652 | -100.0027
4.7688 | -100.0107
4.7725 | -100.0403
4.7762 | -100.0982
4.7798 | -100.1819
4.7835 | -100.2875
4.7872 | -100.3961
4.7908 | -100.5017
4.7945 | -100.5927
4.7982 | -100.6691
4.8018 | -100.7154
4.8055 | -100.7307
4.8092 | -100.7154
4.8128 | -100.6717
4.8165 | -100.5938
4.8202 | -100.4966
4.8238 | -100.3874
4.8275 | -100.2731
4.8312 | -100.162
4.8348 | -100.0704
4.8385 | -99.9958
4.8422 | -99.9382
4.8458 | -99.9035
4.8495 | -99.8863
4.8532 | -99.8827
4.8568 | -99.8891
4.8605 | -99.9014
4.8641 | -99.9168
4.8678 | -99.935
4.8715 | -99.9545
4.8751 | -99.9769
4.8788 | -100.0078
4.8825 | -100.0464
4.8861 | -100.0909
4.8898 | -100.1391
4.8935 | -100.1956
4.8971 | -100.2549
4.9008 | -100.305
4.9045 | -100.3409
4.9081 | -100.3578
4.9118 | -100.3522
4.9155 | -100.3308
4.9191 | -100.2894
4.9228 | -100.2251
4.9265 | -100.1596
4.9301 | -100.1065
4.9338 | -100.0672
4.9375 | -100.0458
4.9411 | -100.0441
4.9448 | -100.0576
4.9485 | -100.0936
4.9521 | -100.1513
4.9558 | -100.2241
4.9595 | -100.3024
4.9631 | -100.3805
4.9668 | -100.4448
4.9704 | -100.4879
4.9741 | -100.519
4.9778 | -100.5381
4.9814 | -100.5411
4.9851 | -100.5311
4.9888 | -100.5051
4.9924 | -100.471
4.9961 | -100.4406
4.9998 | -100.4202
5.0034 | -100.404
5.0071 | -100.3905
5.0108 | -100.393
5.0144 | -100.4037
5.0181 | -100.4175
5.0218 | -100.4315
5.0254 | -100.4355
5.0291 | -100.4224
5.0328 | -100.4049
5.0364 | -100.383
5.0401 | -100.3565
5.0438 | -100.3305
5.0474 | -100.2973
5.0511 | -100.2577
5.0548 | -100.226
5.0584 | -100.1933
5.0621 | -100.1589
5.0658 | -100.1345
5.0694 | -100.1231
5.0731 | -100.1273
5.0767 | -100.1403
5.0804 | -100.1633
5.0841 | -100.2031
5.0877 | -100.2447
5.0914 | -100.2964
5.0951 | -100.3538
5.0987 | -100.3921
5.1024 | -100.4147
5.1061 | -100.4375
5.1097 | -100.4434
5.1134 | -100.4396
5.1171 | -100.4262
5.1207 | -100.4057
5.1244 | -100.3919
5.1281 | -100.3848
5.1317 | -100.3892
5.1354 | -100.412
5.1391 | -100.4488
5.1427 | -100.486
5.1464 | -100.5224
5.1501 | -100.5517
5.1537 | -100.5837
5.1574 | -100.6101
5.1611 | -100.6306
5.1647 | -100.6532
5.1684 | -100.678
5.172 | -100.7196
5.1757 | -100.7859
5.1794 | -100.8701
5.183 | -100.9515
5.1867 | -101.0352
5.1904 | -101.1091
5.194 | -101.1691
5.1977 | -101.2055
5.2014 | -101.2119
5.205 | -101.1885
5.2087 | -101.1387
5.2124 | -101.0627
5.216 | -100.9725
5.2197 | -100.8899
5.2234 | -100.8197
5.227 | -100.7669
5.2307 | -100.7348
5.2344 | -100.7287
5.238 | -100.7474
5.2417 | -100.7973
5.2454 | -100.8647
5.249 | -100.9353
5.2527 | -101.006
5.2564 | -101.0543
5.26 | -101.0808
5.2637 | -101.0836
5.2674 | -101.0611
5.271 | -101.0125
5.2747 | -100.9537
5.2784 | -100.899
5.282 | -100.8524
5.2857 | -100.8521
5.2893 | -100.8321
5.293 | -100.8286
5.2967 | -100.8459
5.3003 | -100.8647
5.304 | -100.8913
5.3077 | -100.9295
5.3113 | -100.9724
5.315 | -101.0144
5.3187 | -101.0449
5.3223 | -101.0529
5.326 | -101.0495
5.3297 | -101.0305
5.3333 | -101.0009
5.337 | -100.9753
5.3407 | -100.9578
5.3443 | -100.9468
5.348 | -100.9516
5.3517 | -100.974
5.3553 | -101.0035
5.359 | -101.028
5.3627 | -101.0463
5.3663 | -101.0539
5.37 | -101.0574
5.3737 | -101.0487
5.3773 | -101.0186
5.381 | -100.9713
5.3847 | -100.9261
5.3883 | -100.8772
5.392 | -100.8289
5.3956 | -100.7882
5.3993 | -100.7674
5.403 | -100.7721
5.4066 | -100.8065
5.4103 | -100.8695
5.414 | -100.9628
5.4176 | -101.0846
5.4213 | -101.2261
5.425 | -101.3737
5.4286 | -101.525
5.4323 | -101.6584
5.436 | -101.772
5.4396 | -101.8495
5.4433 | -101.882
5.447 | -101.8667
5.4506 | -101.8177
5.4543 | -101.735
5.458 | -101.6318
5.4616 | -101.529
5.4653 | -101.4318
5.469 | -101.3409
5.4726 | -101.269
5.4763 | -101.2195
5.48 | -101.1887
5.4836 | -101.1713
5.4873 | -101.1682
5.491 | -101.1666
5.4946 | -101.1635
5.4983 | -101.1661
5.5019 | -101.1633
5.5056 | -101.1487
5.5093 | -101.1304
5.5129 | -101.1022
5.5166 | -101.0651
5.5203 | -101.0204
5.5239 | -100.9885
5.5276 | -100.9706
5.5313 | -100.9621
5.5349 | -100.9572
5.5386 | -100.9641
5.5423 | -100.9858
5.5459 | -101.0095
5.5496 | -101.0281
5.5533 | -101.0499
5.5569 | -101.0716
5.5606 | -101.0869
5.5643 | -101.0994
5.5679 | -101.1068
5.5716 | -101.113
5.5753 | -101.126
5.5789 | -101.1437
5.5826 | -101.1558
5.5863 | -101.1666
5.5899 | -101.1808
5.5936 | -101.1862
5.5973 | -101.1859
5.6009 | -101.1859
5.6046 | -101.1758
5.6082 | -101.1518
5.6119 | -101.1246
5.6156 | -101.0957
5.6192 | -101.0577
5.6229 | -101.0241
5.6266 | -100.9986
5.6302 | -100.9828
5.6339 | -100.9817
5.6376 | -100.99
5.6412 | -101.0128
5.6449 | -101.0539
5.6486 | -101.1035
5.6522 | -101.1661
5.6559 | -101.2313
5.6596 | -101.2912
5.6632 | -101.349
5.6669 | -101.4079
5.6706 | -101.4546
5.6742 | -101.4905
5.6779 | -101.523
5.6816 | -101.5393
5.6852 | -101.5385
5.6889 | -101.5316
5.6926 | -101.5161
5.6962 | -101.4902
5.6999 | -101.4643
5.7036 | -101.4454
5.7072 | -101.4194
5.7109 | -101.3948
5.7145 | -101.3852
5.7182 | -101.3816
5.7219 | -101.3827
5.7255 | -101.3851
5.7292 | -101.3813
5.7329 | -101.3781
5.7365 | -101.3697
5.7402 | -101.3597
5.7439 | -101.3478
5.7475 | -101.3305
5.7512 | -101.3072
5.7549 | -101.29
5.7585 | -101.2749
5.7622 | -101.2538
5.7659 | -101.2355
5.7695 | -101.2178
5.7732 | -101.1957
5.7769 | -101.1682
5.7805 | -101.1444
5.7842 | -101.1213
5.7879 | -101.0996
5.7915 | -101.0797
5.7952 | -101.0643
5.7989 | -101.0486
5.8025 | -101.0461
5.8062 | -101.0579
5.8099 | -101.0844
5.8135 | -101.1401
5.8172 | -101.1814
5.8208 | -101.2327
5.8245 | -101.2869
5.8282 | -101.3354
5.8318 | -101.3853
5.8355 | -101.4329
5.8392 | -101.4801
5.8428 | -101.5222
5.8465 | -101.5588
5.8502 | -101.592
5.8538 | -101.6201
5.8575 | -101.6384
5.8612 | -101.6452
5.8648 | -101.6456
5.8685 | -101.6211
5.8722 | -101.5776
5.8758 | -101.5235
5.8795 | -101.4681
5.8832 | -101.4061
5.8868 | -101.3431
5.8905 | -101.2787
5.8942 | -101.2243
5.8978 | -101.1809
5.9015 | -101.1494
5.9052 | -101.1319
5.9088 | -101.1295
5.9125 | -101.1405
5.9162 | -101.1595
5.9198 | -101.188
5.9235 | -101.2249
5.9271 | -101.2622
5.9308 | -101.3041
5.9345 | -101.3533
5.9381 | -101.4102
5.9418 | -101.4764
5.9455 | -101.5443
5.9491 | -101.6232
5.9528 | -101.7044
5.9565 | -101.7793
5.9601 | -101.8577
5.9638 | -101.9338
5.9675 | -101.9993
5.9711 | -102.0437
5.9748 | -102.0794
5.9785 | -102.0994
5.9821 | -102.1067
5.9858 | -102.0981
5.9895 | -102.0764
5.9931 | -102.0384
5.9968 | -101.9882
6.0005 | -101.931
6.0041 | -101.8649
6.0078 | -101.8019
6.0115 | -101.7407
6.0151 | -101.6754
6.0188 | -101.62
6.0225 | -101.5707
6.0261 | -101.5254
6.0298 | -101.4824
6.0334 | -101.4467
6.0371 | -101.419
6.0408 | -101.3904
6.0444 | -101.3663
6.0481 | -101.3468
6.0518 | -101.3327
6.0554 | -101.3227
6.0591 | -101.3249
6.0628 | -101.3426
6.0664 | -101.3666
6.0701 | -101.4018
6.0738 | -101.4433
6.0774 | -101.483
6.0811 | -101.529
6.0848 | -101.5857
6.0884 | -101.637
6.0921 | -101.672
6.0958 | -101.6986
6.0994 | -101.7132
6.1031 | -101.715
6.1068 | -101.6997
6.1104 | -101.6752
6.1141 | -101.6442
6.1178 | -101.6175
6.1214 | -101.5955
6.1251 | -101.5789
6.1288 | -101.5728
6.1324 | -101.5821
6.1361 | -101.6048
6.1397 | -101.641
6.1434 | -101.6802
6.1471 | -101.7271
6.1507 | -101.7778
6.1544 | -101.8234
6.1581 | -101.8663
6.1617 | -101.9133
6.1654 | -101.9549
6.1691 | -101.9962
6.1727 | -102.023
6.1764 | -102.0445
6.1801 | -102.0682
6.1837 | -102.0892
6.1874 | -102.1029
6.1911 | -102.122
6.1947 | -102.1429
6.1984 | -102.1688
6.2021 | -102.1903
6.2057 | -102.2008
6.2094 | -102.205
6.2131 | -102.2048
6.2167 | -102.1966
6.2204 | -102.1851
6.2241 | -102.1593
6.2277 | -102.1186
6.2314 | -102.0652
6.2351 | -101.999
6.2387 | -101.9239
6.2424 | -101.8446
6.246 | -101.7663
6.2497 | -101.6944
6.2534 | -101.6326
6.257 | -101.5953
6.2607 | -101.5791
6.2644 | -101.5958
6.268 | -101.6308
6.2717 | -101.6818
6.2754 | -101.7541
6.279 | -101.8409
6.2827 | -101.931
6.2864 | -102.0168
6.29 | -102.0981
6.2937 | -102.1708
6.2974 | -102.2361
6.301 | -102.2836
6.3047 | -102.3174
6.3084 | -102.3402
6.312 | -102.3443
6.3157 | -102.3403
6.3194 | -102.3272
6.323 | -102.3081
6.3267 | -102.2935
6.3304 | -102.2775
6.334 | -102.2577
6.3377 | -102.2448
6.3414 | -102.2222
6.345 | -102.1882
6.3487 | -102.1586
6.3523 | -102.1316
6.356 | -102.1088
6.3597 | -102.0881
6.3633 | -102.0797
6.367 | -102.0809
6.3707 | -102.0849
6.3743 | -102.1007
6.378 | -102.1253
6.3817 | -102.1599
6.3853 | -102.2036
6.389 | -102.2517
6.3927 | -102.2987
6.3963 | -102.3355
6.4 | -102.3535
6.4037 | -102.3555
6.4073 | -102.3497
6.411 | -102.3327
6.4147 | -102.302
6.4183 | -102.2546
6.422 | -102.2046
6.4257 | -102.1513
6.4293 | -102.099
6.433 | -102.0517
6.4367 | -102.0171
6.4403 | -101.9905
6.444 | -101.9733
6.4477 | -101.9687
6.4513 | -101.9717
6.455 | -101.9866
6.4586 | -102.0099
6.4623 | -102.0385
6.466 | -102.0715
6.4696 | -102.1068
6.4733 | -102.1409
6.477 | -102.1743
6.4806 | -102.211
6.4843 | -102.2405
6.488 | -102.2709
6.4916 | -102.3092
6.4953 | -102.3458
6.499 | -102.3788
6.5026 | -102.4073
6.5063 | -102.4304
6.51 | -102.4527
6.5136 | -102.4789
6.5173 | -102.5052
6.521 | -102.518
6.5246 | -102.5247
6.5283 | -102.5352
6.532 | -102.539
6.5356 | -102.5396
6.5393 | -102.5277
6.543 | -102.5058
6.5466 | -102.4719
6.5503 | -102.4334
6.554 | -102.3832
6.5576 | -102.3237
6.5613 | -102.2607
6.5649 | -102.2032
6.5686 | -102.1515
6.5723 | -102.1068
6.5759 | -102.0726
6.5796 | -102.0506
6.5833 | -102.0463
6.5869 | -102.0472
6.5906 | -102.0542
6.5943 | -102.0826
6.5979 | -102.121
6.6016 | -102.154
6.6053 | -102.1836
6.6089 | -102.2146
6.6126 | -102.2381
6.6163 | -102.2513
6.6199 | -102.2581
6.6236 | -102.2624
6.6273 | -102.2497
6.6309 | -102.2256
6.6346 | -102.2
6.6383 | -102.1796
6.6419 | -102.1624
6.6456 | -102.1507
6.6493 | -102.1412
6.6529 | -102.1457
6.6566 | -102.1595
6.6603 | -102.179
6.6639 | -102.2032
6.6676 | -102.2328
6.6712 | -102.2647
6.6749 | -102.3053
6.6786 | -102.3529
6.6822 | -102.4033
6.6859 | -102.4662
6.6896 | -102.5258
6.6932 | -102.5733
6.6969 | -102.6198
6.7006 | -102.6579
6.7042 | -102.6781
6.7079 | -102.6788
6.7116 | -102.6632
6.7152 | -102.6339
6.7189 | -102.5962
6.7226 | -102.5414
6.7262 | -102.4822
6.7299 | -102.4227
6.7336 | -102.3574
6.7372 | -102.2997
6.7409 | -102.2497
6.7446 | -102.2031
6.7482 | -102.1698
6.7519 | -102.1492
6.7556 | -102.1396
6.7592 | -102.14
6.7629 | -102.1485
6.7666 | -102.1637
6.7702 | -102.1879
6.7739 | -102.2168
6.7775 | -102.2496
6.7812 | -102.2901
6.7849 | -102.333
6.7885 | -102.3835
6.7922 | -102.4429
6.7959 | -102.5025
6.7995 | -102.5575
6.8032 | -102.6064
6.8069 | -102.6481
6.8105 | -102.6837
6.8142 | -102.7039
6.8179 | -102.7085
6.8215 | -102.7062
6.8252 | -102.7008
6.8289 | -102.686
6.8325 | -102.6674
6.8362 | -102.6411
6.8399 | -102.6129
6.8435 | -102.5904
6.8472 | -102.5697
6.8509 | -102.5477
6.8545 | -102.5251
6.8582 | -102.5076
6.8619 | -102.4935
6.8655 | -102.4777
6.8692 | -102.4617
6.8729 | -102.4344
6.8765 | -102.4106
6.8802 | -102.3847
6.8838 | -102.3583
6.8875 | -102.3394
6.8912 | -102.3256
6.8948 | -102.3156
6.8985 | -102.3142
6.9022 | -102.324
6.9058 | -102.3387
6.9095 | -102.3564
6.9132 | -102.3797
6.9168 | -102.4103
6.9205 | -102.4432
6.9242 | -102.4782
6.9278 | -102.5111
6.9315 | -102.5528
6.9352 | -102.5931
6.9388 | -102.6384
6.9425 | -102.6887
6.9462 | -102.7413
6.9498 | -102.7898
6.9535 | -102.8396
6.9572 | -102.8891
6.9608 | -102.9332
6.9645 | -102.9718
6.9682 | -103.0064
6.9718 | -103.024
6.9755 | -103.0367
6.9792 | -103.0398
6.9828 | -103.0271
6.9865 | -103.0034
6.9901 | -102.9741
6.9938 | -102.9313
6.9975 | -102.8861
7.0011 | -102.8356
7.0048 | -102.7813
7.0085 | -102.7317
7.0121 | -102.6858
7.0158 | -102.6433
7.0195 | -102.6013
7.0231 | -102.5619
7.0268 | -102.5193
7.0305 | -102.4827
7.0341 | -102.4533
7.0378 | -102.423
7.0415 | -102.397
7.0451 | -102.3736
7.0488 | -102.3546
7.0525 | -102.3407
7.0561 | -102.3347
7.0598 | -102.3328
7.0635 | -102.334
7.0671 | -102.3339
7.0708 | -102.3356
7.0745 | -102.3439
7.0781 | -102.3597
7.0818 | -102.3775
7.0855 | -102.3995
7.0891 | -102.4264
7.0928 | -102.4586
7.0964 | -102.4914
7.1001 | -102.5358
7.1038 | -102.5859
7.1074 | -102.643
7.1111 | -102.7036
7.1148 | -102.7556
7.1184 | -102.8036
7.1221 | -102.8546
7.1258 | -102.9017
7.1294 | -102.9418
7.1331 | -102.9778
7.1368 | -103.0055
7.1404 | -103.0292
7.1441 | -103.0525
7.1478 | -103.0709
7.1514 | -103.0828
7.1551 | -103.0902
7.1588 | -103.0936
7.1624 | -103.0899
7.1661 | -103.0844
7.1698 | -103.0755
7.1734 | -103.0624
7.1771 | -103.0427
7.1808 | -103.0147
7.1844 | -102.9841
7.1881 | -102.9462
7.1918 | -102.9075
7.1954 | -102.8722
7.1991 | -102.8392
7.2027 | -102.8115
7.2064 | -102.7836
7.2101 | -102.7585
7.2137 | -102.7415
7.2174 | -102.7303
7.2211 | -102.7222
7.2247 | -102.7083
7.2284 | -102.6905
7.2321 | -102.6788
7.2357 | -102.6642
7.2394 | -102.6436
7.2431 | -102.6203
7.2467 | -102.6002
7.2504 | -102.5808
7.2541 | -102.5652
7.2577 | -102.5549
7.2614 | -102.5478
7.2651 | -102.5396
7.2687 | -102.5324
7.2724 | -102.531
7.2761 | -102.5388
7.2797 | -102.549
7.2834 | -102.5658
7.2871 | -102.5944
7.2907 | -102.6249
7.2944 | -102.6607
7.2981 | -102.6971
7.3017 | -102.7296
7.3054 | -102.7538
7.309 | -102.7726
7.3127 | -102.7866
7.3164 | -102.7902
7.32 | -102.7827
7.3237 | -102.7604
7.3274 | -102.731
7.331 | -102.6982
7.3347 | -102.6532
7.3384 | -102.6014
7.342 | -102.5513
7.3457 | -102.4995
7.3494 | -102.4488
7.353 | -102.4036
7.3567 | -102.3712
7.3604 | -102.3558
7.364 | -102.3506
7.3677 | -102.356
7.3714 | -102.38
7.375 | -102.4124
7.3787 | -102.4513
7.3824 | -102.5003
7.386 | -102.5575
7.3897 | -102.6184
7.3934 | -102.6823
7.397 | -102.7506
7.4007 | -102.8043
7.4044 | -102.8669
7.408 | -102.9269
7.4117 | -102.9805
7.4153 | -103.0282
7.419 | -103.0672
7.4227 | -103.0972
7.4263 | -103.1227
7.43 | -103.1349
7.4337 | -103.1381
7.4373 | -103.134
7.441 | -103.12
7.4447 | -103.1032
7.4483 | -103.0814
7.452 | -103.0619
7.4557 | -103.042
7.4593 | -103.0148
7.463 | -102.976
7.4667 | -102.9359
7.4703 | -102.895
7.474 | -102.8591
7.4777 | -102.8208
7.4813 | -102.7886
7.485 | -102.7632
7.4887 | -102.7298
7.4923 | -102.7015
7.496 | -102.6837
7.4997 | -102.6686
7.5033 | -102.6542
7.507 | -102.6452
7.5107 | -102.641
7.5143 | -102.641
7.518 | -102.6383
7.5216 | -102.6408
7.5253 | -102.6476
7.529 | -102.659
7.5326 | -102.6783
7.5363 | -102.7073
7.54 | -102.7439
7.5436 | -102.7925
7.5473 | -102.8531
7.551 | -102.9225
7.5546 | -103.0003
7.5583 | -103.0856
7.562 | -103.1762
7.5656 | -103.2631
7.5693 | -103.3517
7.573 | -103.4355
7.5766 | -103.5144
7.5803 | -103.5785
7.584 | -103.6252
7.5876 | -103.6536
7.5913 | -103.6719
7.595 | -103.6678
7.5986 | -103.6521
7.6023 | -103.6336
7.606 | -103.6002
7.6096 | -103.5585
7.6133 | -103.5173
7.617 | -103.4737
7.6206 | -103.4303
7.6243 | -103.3875
7.628 | -103.3472
7.6316 | -103.3088
7.6353 | -103.2745
7.6389 | -103.2438
7.6426 | -103.2137
7.6463 | -103.1838
7.6499 | -103.1546
7.6536 | -103.1281
7.6573 | -103.106
7.6609 | -103.083
7.6646 | -103.0589
7.6683 | -103.0423
7.6719 | -103.0284
7.6756 | -103.0244
7.6793 | -103.0318
7.6829 | -103.0474
7.6866 | -103.0724
7.6903 | -103.1058
7.6939 | -103.151
7.6976 | -103.2067
7.7013 | -103.2691
7.7049 | -103.3352
7.7086 | -103.404
7.7123 | -103.4746
7.7159 | -103.5428
7.7196 | -103.6013
7.7233 | -103.6512
7.7269 | -103.69
7.7306 | -103.7098
7.7343 | -103.7179
7.7379 | -103.7163
7.7416 | -103.7039
7.7452 | -103.6843
7.7489 | -103.6546
7.7526 | -103.6205
7.7562 | -103.5848
7.7599 | -103.5525
7.7636 | -103.5217
7.7672 | -103.4914
7.7709 | -103.4597
7.7746 | -103.4344
7.7782 | -103.4126
7.7819 | -103.3885
7.7856 | -103.3649
7.7892 | -103.3397
7.7929 | -103.3148
7.7966 | -103.2908
7.8002 | -103.267
7.8039 | -103.2465
7.8076 | -103.2223
7.8112 | -103.1997
7.8149 | -103.1826
7.8186 | -103.1713
7.8222 | -103.1587
7.8259 | -103.1451
7.8296 | -103.1321
7.8332 | -103.123
7.8369 | -103.1232
7.8406 | -103.1279
7.8442 | -103.139
7.8479 | -103.1537
7.8515 | -103.1678
7.8552 | -103.1806
7.8589 | -103.1988
7.8625 | -103.2202
7.8662 | -103.2428
7.8699 | -103.2684
7.8735 | -103.2939
7.8772 | -103.3212
7.8809 | -103.3547
7.8845 | -103.3948
7.8882 | -103.4364
7.8919 | -103.4811
7.8955 | -103.53
7.8992 | -103.583
7.9029 | -103.6346
7.9065 | -103.6867
7.9102 | -103.7346
7.9139 | -103.7779
7.9175 | -103.8124
7.9212 | -103.8396
7.9249 | -103.857
7.9285 | -103.8707
7.9322 | -103.8528
7.9359 | -103.8526
7.9395 | -103.8489
7.9432 | -103.838
7.9469 | -103.8261
7.9505 | -103.8083
7.9542 | -103.7762
7.9578 | -103.744
7.9615 | -103.717
7.9652 | -103.6883
7.9688 | -103.6622
7.9725 | -103.6383
7.9762 | -103.6099
7.9798 | -103.5854
7.9835 | -103.5626
7.9872 | -103.5399
7.9908 | -103.5175
7.9945 | -103.4997
7.9982 | -103.4861
8.0018 | -103.4684
8.0055 | -103.452
8.0092 | -103.4369
8.0128 | -103.4275
8.0165 | -103.4209
8.0202 | -103.4186
8.0238 | -103.4157
8.0275 | -103.4118
8.0312 | -103.4091
8.0348 | -103.4059
8.0385 | -103.4037
8.0422 | -103.3977
8.0458 | -103.3872
8.0495 | -103.3737
8.0532 | -103.3572
8.0568 | -103.3359
8.0605 | -103.3166
8.0641 | -103.2983
8.0678 | -103.2793
8.0715 | -103.2652
8.0751 | -103.2523
8.0788 | -103.2387
8.0825 | -103.2287
8.0861 | -103.2219
8.0898 | -103.2189
8.0935 | -103.2243
8.0971 | -103.2368
8.1008 | -103.2496
8.1045 | -103.2682
8.1081 | -103.2922
8.1118 | -103.3196
8.1155 | -103.3508
8.1191 | -103.3786
8.1228 | -103.4162
8.1265 | -103.4512
8.1301 | -103.4844
8.1338 | -103.5121
8.1375 | -103.5342
8.1411 | -103.5508
8.1448 | -103.5631
8.1485 | -103.5712
8.1521 | -103.5686
8.1558 | -103.5607
8.1594 | -103.5469
8.1631 | -103.5279
8.1668 | -103.5071
8.1704 | -103.4931
8.1741 | -103.4815
8.1778 | -103.4702
8.1814 | -103.4626
8.1851 | -103.4578
8.1888 | -103.463
8.1924 | -103.4735
8.1961 | -103.4921
8.1998 | -103.5117
8.2034 | -103.5269
8.2071 | -103.5449
8.2108 | -103.5591
8.2144 | -103.5675
8.2181 | -103.5687
8.2218 | -103.5672
8.2254 | -103.5607
8.2291 | -103.5495
8.2328 | -103.5352
8.2364 | -103.5166
8.2401 | -103.4989
8.2438 | -103.4784
8.2474 | -103.4613
8.2511 | -103.4442
8.2548 | -103.4302
8.2584 | -103.4256
8.2621 | -103.4286
8.2658 | -103.4374
8.2694 | -103.449
8.2731 | -103.4666
8.2767 | -103.4862
8.2804 | -103.5082
8.2841 | -103.5333
8.2877 | -103.5539
8.2914 | -103.5675
8.2951 | -103.5747
8.2987 | -103.578
8.3024 | -103.5757
8.3061 | -103.5696
8.3097 | -103.555
8.3134 | -103.5367
8.3171 | -103.5162
8.3207 | -103.4956
8.3244 | -103.4794
8.3281 | -103.4622
8.3317 | -103.4497
8.3354 | -103.4413
8.3391 | -103.4368
8.3427 | -103.4326
8.3464 | -103.4351
8.3501 | -103.4403
8.3537 | -103.4493
8.3574 | -103.465
8.3611 | -103.4838
8.3647 | -103.5067
8.3684 | -103.5361
8.372 | -103.5692
8.3757 | -103.6024
8.3794 | -103.6304
8.383 | -103.6559
8.3867 | -103.6818
8.3904 | -103.713
8.394 | -103.7453
8.3977 | -103.7752
8.4014 | -103.7986
8.405 | -103.8193
8.4087 | -103.8415
8.4124 | -103.8625
8.416 | -103.8784
8.4197 | -103.8899
8.4234 | -103.9013
8.427 | -103.9102
8.4307 | -103.916
8.4344 | -103.9224
8.438 | -103.927
8.4417 | -103.937
8.4454 | -103.9472
8.449 | -103.9599
8.4527 | -103.9738
8.4564 | -103.9865
8.46 | -103.9875
8.4637 | -103.9839
8.4674 | -103.9783
8.471 | -103.968
8.4747 | -103.9537
8.4784 | -103.9337
8.482 | -103.9071
8.4857 | -103.8755
8.4893 | -103.8368
8.493 | -103.7923
8.4967 | -103.7447
8.5003 | -103.6937
8.504 | -103.6403
8.5077 | -103.5876
8.5113 | -103.5365
8.515 | -103.4867
8.5187 | -103.4422
8.5223 | -103.4057
8.526 | -103.3818
8.5297 | -103.364
8.5333 | -103.3522
8.537 | -103.351
8.5407 | -103.3588
8.5443 | -103.3738
8.548 | -103.39
8.5517 | -103.4115
8.5553 | -103.4377
8.559 | -103.4635
8.5627 | -103.4909
8.5663 | -103.5116
8.57 | -103.5224
8.5737 | -103.5284
8.5773 | -103.5279
8.581 | -103.5206
8.5847 | -103.5118
8.5883 | -103.5015
8.592 | -103.4886
8.5956 | -103.4691
8.5993 | -103.4475
8.603 | -103.4253
8.6066 | -103.4028
8.6103 | -103.3798
8.614 | -103.3593
8.6176 | -103.3417
8.6213 | -103.328
8.625 | -103.3146
8.6286 | -103.3059
8.6323 | -103.2991
8.636 | -103.2907
8.6396 | -103.2812
8.6433 | -103.2712
8.647 | -103.259
8.6506 | -103.2444
8.6543 | -103.2276
8.658 | -103.2122
8.6616 | -103.2003
8.6653 | -103.1867
8.669 | -103.1765
8.6726 | -103.1667
8.6763 | -103.1615
8.68 | -103.1613
8.6836 | -103.1599
8.6873 | -103.1659
8.691 | -103.1767
8.6946 | -103.1906
8.6983 | -103.2129
8.7019 | -103.2378
8.7056 | -103.2692
8.7093 | -103.3022
8.7129 | -103.3356
8.7166 | -103.366
8.7203 | -103.4044
8.7239 | -103.4478
8.7276 | -103.4907
8.7313 | -103.5333
8.7349 | -103.5764
8.7386 | -103.6152
8.7423 | -103.6495
8.7459 | -103.6827
8.7496 | -103.7099
8.7533 | -103.7355
8.7569 | -103.7544
8.7606 | -103.7744
8.7643 | -103.7904
8.7679 | -103.7976
8.7716 | -103.8056
8.7753 | -103.8111
8.7789 | -103.8176
8.7826 | -103.8222
8.7863 | -103.825
8.7899 | -103.8219
8.7936 | -103.8198
8.7973 | -103.8169
8.8009 | -103.8168
8.8046 | -103.8184
8.8082 | -103.8213
8.8119 | -103.8283
8.8156 | -103.8343
8.8192 | -103.8463
8.8229 | -103.8629
8.8266 | -103.8768
8.8302 | -103.8938
8.8339 | -103.9066
8.8376 | -103.9168
8.8412 | -103.9255
8.8449 | -103.9363
8.8486 | -103.9428
8.8522 | -103.9433
8.8559 | -103.9399
8.8596 | -103.9318
8.8632 | -103.917
8.8669 | -103.8986
8.8706 | -103.8728
8.8742 | -103.838
8.8779 | -103.8023
8.8816 | -103.7664
8.8852 | -103.7271
8.8889 | -103.6847
8.8926 | -103.6465
8.8962 | -103.6193
8.8999 | -103.5926
8.9036 | -103.5736
8.9072 | -103.5605
8.9109 | -103.5501
8.9145 | -103.5489
8.9182 | -103.558
8.9219 | -103.5737
8.9255 | -103.5939
8.9292 | -103.6178
8.9329 | -103.6461
8.9365 | -103.6747
8.9402 | -103.7107
8.9439 | -103.7505
8.9475 | -103.7932
8.9512 | -103.8421
8.9549 | -103.8938
8.9585 | -103.9458
8.9622 | -104.0021
8.9659 | -104.0593
8.9695 | -104.1094
8.9732 | -104.1462
8.9769 | -104.1736
8.9805 | -104.1947
8.9842 | -104.2093
8.9879 | -104.1995
8.9915 | -104.205
8.9952 | -104.2091
8.9989 | -104.2081
9.0025 | -104.2063
9.0062 | -104.2036
9.0099 | -104.1958
9.0135 | -104.1821
9.0172 | -104.1755
9.0208 | -104.174
9.0245 | -104.1737
9.0282 | -104.1715
9.0318 | -104.1751
9.0355 | -104.1763
9.0392 | -104.1752
9.0428 | -104.1813
9.0465 | -104.1826
9.0502 | -104.1792
9.0538 | -104.1758
9.0575 | -104.1707
9.0612 | -104.1649
9.0648 | -104.1579
9.0685 | -104.1485
9.0722 | -104.1313
9.0758 | -104.1148
9.0795 | -104.0965
9.0832 | -104.0758
9.0868 | -104.0561
9.0905 | -104.0385
9.0942 | -104.0188
9.0978 | -103.9991
9.1015 | -103.9851
9.1052 | -103.9782
9.1088 | -103.9725
9.1125 | -103.9707
9.1162 | -103.9691
9.1198 | -103.9672
9.1235 | -103.9728
9.1271 | -103.9811
9.1308 | -103.9902
9.1345 | -104.0027
9.1381 | -104.0196
9.1418 | -104.0359
9.1455 | -104.0533
9.1491 | -104.0739
9.1528 | -104.0906
9.1565 | -104.1032
9.1601 | -104.1129
9.1638 | -104.1232
9.1675 | -104.1297
9.1711 | -104.1317
9.1748 | -104.1321
9.1785 | -104.1246
9.1821 | -104.116
9.1858 | -104.1107
9.1895 | -104.1018
9.1931 | -104.0922
9.1968 | -104.0831
9.2005 | -104.0761
9.2041 | -104.0735
9.2078 | -104.0698
9.2115 | -104.0632
9.2151 | -104.0603
9.2188 | -104.0618
9.2225 | -104.0663
9.2261 | -104.0706
9.2298 | -104.0815
9.2334 | -104.0974
9.2371 | -104.1133
9.2408 | -104.1311
9.2444 | -104.1552
9.2481 | -104.181
9.2518 | -104.2109
9.2554 | -104.2448
9.2591 | -104.2779
9.2628 | -104.3137
9.2664 | -104.3471
9.2701 | -104.3796
9.2738 | -104.4117
9.2774 | -104.4448
9.2811 | -104.4756
9.2848 | -104.4973
9.2884 | -104.5142
9.2921 | -104.5255
9.2958 | -104.5383
9.2994 | -104.5439
9.3031 | -104.5414
9.3068 | -104.5401
9.3104 | -104.534
9.3141 | -104.5258
9.3178 | -104.5105
9.3214 | -104.4831
9.3251 | -104.4583
9.3288 | -104.4301
9.3324 | -104.3996
9.3361 | -104.3656
9.3397 | -104.3336
9.3434 | -104.3038
9.3471 | -104.274
9.3507 | -104.2349
9.3544 | -104.1958
9.3581 | -104.1616
9.3617 | -104.1301
9.3654 | -104.097
9.3691 | -104.0665
9.3727 | -104.04
9.3764 | -104.0163
9.3801 | -103.9939
9.3837 | -103.9811
9.3874 | -103.9726
9.3911 | -103.9638
9.3947 | -103.9634
9.3984 | -103.9669
9.4021 | -103.9758
9.4057 | -103.9919
9.4094 | -104.0099
9.4131 | -104.0356
9.4167 | -104.066
9.4204 | -104.1025
9.4241 | -104.1385
9.4277 | -104.1788
9.4314 | -104.2168
9.4351 | -104.2517
9.4387 | -104.2908
9.4424 | -104.3271
9.446 | -104.3608
9.4497 | -104.39
9.4534 | -104.4142
9.457 | -104.4354
9.4607 | -104.447
9.4644 | -104.4522
9.468 | -104.4517
9.4717 | -104.4451
9.4754 | -104.43
9.479 | -104.4136
9.4827 | -104.3939
9.4864 | -104.3753
9.49 | -104.3497
9.4937 | -104.3133
9.4974 | -104.2743
9.501 | -104.2396
9.5047 | -104.2033
9.5084 | -104.1657
9.512 | -104.1252
9.5157 | -104.0794
9.5194 | -104.0394
9.523 | -103.9975
9.5267 | -103.9594
9.5304 | -103.9271
9.534 | -103.8896
9.5377 | -103.8572
9.5414 | -103.8319
9.545 | -103.8083
9.5487 | -103.7898
9.5523 | -103.7718
9.556 | -103.7581
9.5597 | -103.7518
9.5633 | -103.7502
9.567 | -103.7533
9.5707 | -103.7582
9.5743 | -103.7673
9.578 | -103.777
9.5817 | -103.792
9.5853 | -103.8071
9.589 | -103.8217
9.5927 | -103.8401
9.5963 | -103.8568
9.6 | -103.8729
9.6037 | -103.891
9.6073 | -103.9038
9.611 | -103.9144
9.6147 | -103.9261
9.6183 | -103.9358
9.622 | -103.9463
9.6257 | -103.9551
9.6293 | -103.966
9.633 | -103.9819
9.6367 | -103.9969
9.6403 | -104.0066
9.644 | -104.0139
9.6477 | -104.0219
9.6513 | -104.0333
9.655 | -104.0429
9.6586 | -104.0563
9.6623 | -104.0661
9.666 | -104.0727
9.6696 | -104.0763
9.6733 | -104.0815
9.677 | -104.0906
9.6806 | -104.094
9.6843 | -104.0962
9.688 | -104.1
9.6916 | -104.1058
9.6953 | -104.1177
9.699 | -104.1291
9.7026 | -104.1413
9.7063 | -104.1538
9.71 | -104.1636
9.7136 | -104.1773
9.7173 | -104.1958
9.721 | -104.2136
9.7246 | -104.2282
9.7283 | -104.2407
9.732 | -104.2572
9.7356 | -104.2738
9.7393 | -104.2875
9.743 | -104.295
9.7466 | -104.3002
9.7503 | -104.3082
9.754 | -104.3113
9.7576 | -104.3078
9.7613 | -104.3005
9.7649 | -104.2862
9.7686 | -104.2703
9.7723 | -104.2572
9.7759 | -104.2385
9.7796 | -104.2199
9.7833 | -104.2065
9.7869 | -104.1934
9.7906 | -104.1788
9.7943 | -104.1625
9.7979 | -104.1487
9.8016 | -104.1388
9.8053 | -104.1305
9.8089 | -104.1275
9.8126 | -104.1263
9.8163 | -104.1249
9.8199 | -104.1251
9.8236 | -104.1281
9.8273 | -104.1308
9.8309 | -104.1329
9.8346 | -104.136
9.8383 | -104.1338
9.8419 | -104.1363
9.8456 | -104.1385
9.8493 | -104.1403
9.8529 | -104.1475
9.8566 | -104.1538
9.8603 | -104.1607
9.8639 | -104.1653
9.8676 | -104.1698
9.8712 | -104.1781
9.8749 | -104.1817
9.8786 | -104.1826
9.8822 | -104.1854
9.8859 | -104.1855
9.8896 | -104.1849
9.8932 | -104.1839
9.8969 | -104.1868
9.9006 | -104.1916
9.9042 | -104.1955
9.9079 | -104.201
9.9116 | -104.2106
9.9152 | -104.2137
9.9189 | -104.2177
9.9226 | -104.2233
9.9262 | -104.2291
9.9299 | -104.2377
9.9336 | -104.2494
9.9372 | -104.2614
9.9409 | -104.2748
9.9446 | -104.2861
9.9482 | -104.2956
9.9519 | -104.3044
9.9556 | -104.3163
9.9592 | -104.3286
9.9629 | -104.3429
9.9666 | -104.3566
9.9702 | -104.3707
9.9739 | -104.3879
9.9775 | -104.4039
9.9812 | -104.4256
9.9849 | -104.4422
9.9885 | -104.4587
9.9922 | -104.4804
9.9959 | -104.4984
9.9995 | -104.5137
10.0032 | -104.5269
10.0069 | -104.5418
10.0105 | -104.5546
10.0142 | -104.5668
10.0179 | -104.5788
10.0215 | -104.591
10.0252 | -104.5994
10.0289 | -104.6068
10.0325 | -104.6077
10.0362 | -104.6031
10.0399 | -104.5898
10.0435 | -104.5708
10.0472 | -104.539
10.0509 | -104.5167
10.0545 | -104.495
10.0582 | -104.4672
10.0619 | -104.4455
10.0655 | -104.4213
10.0692 | -104.3914
10.0729 | -104.3587
10.0765 | -104.3293
10.0802 | -104.2975
10.0838 | -104.2697
10.0875 | -104.2407
10.0912 | -104.2118
10.0948 | -104.1887
10.0985 | -104.1643
10.1022 | -104.1383
10.1058 | -104.1133
10.1095 | -104.0895
10.1132 | -104.0702
10.1168 | -104.0527
10.1205 | -104.038
10.1242 | -104.0292
10.1278 | -104.0217
10.1315 | -104.0175
10.1352 | -104.0174
10.1388 | -104.0205
10.1425 | -104.0284
10.1462 | -104.0427
10.1498 | -104.06
10.1535 | -104.0812
10.1572 | -104.1074
10.1608 | -104.1336
10.1645 | -104.1662
10.1682 | -104.1972
10.1718 | -104.2289
10.1755 | -104.2642
10.1792 | -104.3015
10.1828 | -104.3388
10.1865 | -104.3811
10.1901 | -104.4242
10.1938 | -104.4641
10.1975 | -104.4977
10.2011 | -104.5293
10.2048 | -104.5598
10.2085 | -104.5874
10.2121 | -104.6058
10.2158 | -104.6262
10.2195 | -104.6458
10.2231 | -104.6649
10.2268 | -104.6719
10.2305 | -104.6785
10.2341 | -104.6774
10.2378 | -104.6751
10.2415 | -104.673
10.2451 | -104.6703
10.2488 | -104.67
10.2525 | -104.6698
10.2561 | -104.6669
10.2598 | -104.6609
10.2635 | -104.6559
10.2671 | -104.6538
10.2708 | -104.6455
10.2745 | -104.6389
10.2781 | -104.6371
10.2818 | -104.6397
10.2855 | -104.6398
10.2891 | -104.6361
10.2928 | -104.6321
10.2964 | -104.6316
10.3001 | -104.6317
10.3038 | -104.6344
10.3074 | -104.6371
10.3111 | -104.6362
10.3148 | -104.6325
10.3184 | -104.6373
10.3221 | -104.6427
10.3258 | -104.6423
10.3294 | -104.639
10.3331 | -104.6381
10.3368 | -104.6343
10.3404 | -104.6265
10.3441 | -104.6172
10.3478 | -104.6155
10.3514 | -104.6093
10.3551 | -104.5961
10.3588 | -104.5849
10.3624 | -104.5743
10.3661 | -104.5594
10.3698 | -104.5428
10.3734 | -104.5271
10.3771 | -104.5155
10.3808 | -104.4995
10.3844 | -104.4843
10.3881 | -104.4697
10.3918 | -104.4539
10.3954 | -104.4366
10.3991 | -104.4223
10.4027 | -104.4113
10.4064 | -104.4064
10.4101 | -104.4055
10.4137 | -104.4065
10.4174 | -104.4077
10.4211 | -104.4177
10.4247 | -104.4286
10.4284 | -104.4385
10.4321 | -104.453
10.4357 | -104.4731
10.4394 | -104.4943
10.4431 | -104.5149
10.4467 | -104.5367
10.4504 | -104.5579
10.4541 | -104.5789
10.4577 | -104.6005
10.4614 | -104.6243
10.4651 | -104.6419
10.4687 | -104.6564
10.4724 | -104.6714
10.4761 | -104.6877
10.4797 | -104.7024
10.4834 | -104.7133
10.4871 | -104.726
10.4907 | -104.7338
10.4944 | -104.7363
10.4981 | -104.7338
10.5017 | -104.7339
10.5054 | -104.7355
10.509 | -104.7316
10.5127 | -104.7256
10.5164 | -104.7172
10.52 | -104.7107
10.5237 | -104.7035
10.5274 | -104.6952
10.531 | -104.687
10.5347 | -104.6769
10.5384 | -104.6648
10.542 | -104.6563
10.5457 | -104.6525
10.5494 | -104.6444
10.553 | -104.6362
10.5567 | -104.6295
10.5604 | -104.6221
10.564 | -104.6139
10.5677 | -104.6074
10.5714 | -104.6028
10.575 | -104.6073
10.5787 | -104.6087
10.5824 | -104.613
10.586 | -104.6157
10.5897 | -104.6181
10.5934 | -104.6269
10.597 | -104.6379
10.6007 | -104.6429
10.6044 | -104.6507
10.608 | -104.6588
10.6117 | -104.6698
10.6153 | -104.6825
10.619 | -104.6997
10.6227 | -104.7117
10.6263 | -104.7237
10.63 | -104.7374
10.6337 | -104.7466
10.6373 | -104.7545
10.641 | -104.7618
10.6447 | -104.7738
10.6483 | -104.7863
10.652 | -104.7973
10.6557 | -104.8024
10.6593 | -104.8024
10.663 | -104.8031
10.6667 | -104.8056
10.6703 | -104.8065
10.674 | -104.8043
10.6777 | -104.7991
10.6813 | -104.7983
10.685 | -104.7997
10.6887 | -104.8008
10.6923 | -104.799
10.696 | -104.7943
10.6997 | -104.79
10.7033 | -104.7875
10.707 | -104.7839
10.7107 | -104.7783
10.7143 | -104.7787
10.718 | -104.7802
10.7216 | -104.782
10.7253 | -104.7834
10.729 | -104.7838
10.7326 | -104.7856
10.7363 | -104.792
10.74 | -104.8013
10.7436 | -104.8127
10.7473 | -104.827
10.751 | -104.8415
10.7546 | -104.8574
10.7583 | -104.8767
10.762 | -104.8941
10.7656 | -104.9125
10.7693 | -104.9291
10.773 | -104.9449
10.7766 | -104.9577
10.7803 | -104.9702
10.784 | -104.9829
10.7876 | -104.9898
10.7913 | -104.9969
10.795 | -105.0024
10.7986 | -105.0055
10.8023 | -105.0065
10.806 | -105.0051
10.8096 | -104.9987
10.8133 | -104.9944
10.817 | -104.9907
10.8206 | -104.9843
10.8243 | -104.9749
10.8279 | -104.9668
10.8316 | -104.9572
10.8353 | -104.949
10.8389 | -104.9421
10.8426 | -104.9349
10.8463 | -104.9266
10.8499 | -104.9222
10.8536 | -104.9192
10.8573 | -104.9141
10.8609 | -104.9068
10.8646 | -104.9009
10.8683 | -104.8969
10.8719 | -104.889
10.8756 | -104.8788
10.8793 | -104.8681
10.8829 | -104.8565
10.8866 | -104.8444
10.8903 | -104.832
10.8939 | -104.8153
10.8976 | -104.8013
10.9013 | -104.7873
10.9049 | -104.7718
10.9086 | -104.759
10.9123 | -104.7489
10.9159 | -104.7419
10.9196 | -104.7371
10.9233 | -104.7314
10.9269 | -104.7269
10.9306 | -104.7246
10.9343 | -104.7245
10.9379 | -104.7251
10.9416 | -104.7241
10.9452 | -104.7297
10.9489 | -104.7351
10.9526 | -104.7402
10.9562 | -104.7462
10.9599 | -104.7525
10.9636 | -104.7603
10.9672 | -104.7702
10.9709 | -104.7774
10.9746 | -104.7856
10.9782 | -104.7908
10.9819 | -104.794
10.9856 | -104.7957
10.9892 | -104.7991
10.9929 | -104.7964
10.9966 | -104.7924
11.0002 | -104.7898
11.0039 | -104.7915
11.0076 | -104.7913
11.0112 | -104.7895
11.0149 | -104.7909
11.0186 | -104.7949
11.0222 | -104.7953
11.0259 | -104.7951
11.0296 | -104.7975
11.0332 | -104.8055
11.0369 | -104.8116
11.0405 | -104.8165
11.0442 | -104.8277
11.0479 | -104.8398
11.0515 | -104.8529
11.0552 | -104.8621
11.0589 | -104.8735
11.0625 | -104.8877
11.0662 | -104.896
11.0699 | -104.9047
11.0735 | -104.9123
11.0772 | -104.9183
11.0809 | -104.924
11.0845 | -104.93
11.0882 | -104.9353
11.0919 | -104.9415
11.0955 | -104.945
11.0992 | -104.9489
11.1029 | -104.9458
11.1065 | -104.9448
11.1102 | -104.9455
11.1139 | -104.9421
11.1175 | -104.9328
11.1212 | -104.9265
11.1249 | -104.9226
11.1285 | -104.9181
11.1322 | -104.9145
11.1359 | -104.9115
11.1395 | -104.9112
11.1432 | -104.9087
11.1469 | -104.9079
11.1505 | -104.9088
11.1542 | -104.9138
11.1578 | -104.9175
11.1615 | -104.9183
11.1652 | -104.9215
11.1688 | -104.9268
11.1725 | -104.9317
11.1762 | -104.9379
11.1798 | -104.949
11.1835 | -104.961
11.1872 | -104.9713
11.1908 | -104.9781
11.1945 | -104.9839
11.1982 | -104.9908
11.2018 | -104.9944
11.2055 | -104.9922
11.2092 | -104.9905
11.2128 | -104.9873
11.2165 | -104.9827
11.2202 | -104.9789
11.2238 | -104.9698
11.2275 | -104.9571
11.2312 | -104.9462
11.2348 | -104.9344
11.2385 | -104.9218
11.2422 | -104.9071
11.2458 | -104.8903
11.2495 | -104.8717
11.2531 | -104.8515
11.2568 | -104.834
11.2605 | -104.8173
11.2641 | -104.8075
11.2678 | -104.8005
11.2715 | -104.7945
11.2751 | -104.7912
11.2788 | -104.7939
11.2825 | -104.8008
11.2861 | -104.8064
11.2898 | -104.8122
11.2935 | -104.8239
11.2971 | -104.8366
11.3008 | -104.849
11.3045 | -104.8598
11.3081 | -104.8722
11.3118 | -104.8857
11.3155 | -104.8994
11.3191 | -104.9147
11.3228 | -104.9298
11.3265 | -104.9436
11.3301 | -104.9503
11.3338 | -104.9536
11.3375 | -104.9557
11.3411 | -104.958
11.3448 | -104.9549
11.3485 | -104.9492
11.3521 | -104.9425
11.3558 | -104.9338
11.3595 | -104.9214
11.3631 | -104.9098
11.3668 | -104.9006
11.3704 | -104.8833
11.3741 | -104.8653
11.3778 | -104.85
11.3814 | -104.8397
11.3851 | -104.8287
11.3888 | -104.8225
11.3924 | -104.8214
11.3961 | -104.8156
11.3998 | -104.8158
11.4034 | -104.8187
11.4071 | -104.8252
11.4108 | -104.8386
11.4144 | -104.8562
11.4181 | -104.873
11.4218 | -104.888
11.4254 | -104.9057
11.4291 | -104.9215
11.4328 | -104.9405
11.4364 | -104.9633
11.4401 | -104.9855
11.4438 | -105.0031
11.4474 | -105.0217
11.4511 | -105.0399
11.4548 | -105.0541
11.4584 | -105.0633
11.4621 | -105.0735
11.4657 | -105.0825
11.4694 | -105.091
11.4731 | -105.0942
11.4767 | -105.0899
11.4804 | -105.0797
11.4841 | -105.0693
11.4877 | -105.0535
11.4914 | -105.0362
11.4951 | -105.0147
11.4987 | -104.9924
11.5024 | -104.9645
11.5061 | -104.9346
11.5097 | -104.9072
11.5134 | -104.8781
11.5171 | -104.845
11.5207 | -104.8154
11.5244 | -104.7871
11.5281 | -104.7578
11.5317 | -104.7306
11.5354 | -104.7097
11.5391 | -104.6849
11.5427 | -104.6617
11.5464 | -104.6385
11.5501 | -104.6162
11.5537 | -104.5935
11.5574 | -104.5751
11.5611 | -104.5624
11.5647 | -104.5502
11.5684 | -104.5423
11.5721 | -104.5377
11.5757 | -104.5345
11.5794 | -104.5363
11.583 | -104.5441
11.5867 | -104.5596
11.5904 | -104.5726
11.594 | -104.5882
11.5977 | -104.6095
11.6014 | -104.6317
11.605 | -104.6565
11.6087 | -104.6886
11.6124 | -104.72
11.616 | -104.7513
11.6197 | -104.7847
11.6234 | -104.8172
11.627 | -104.8467
11.6307 | -104.8757
11.6344 | -104.9039
11.638 | -104.9288
11.6417 | -104.9542
11.6454 | -104.9789
11.649 | -105.0032
11.6527 | -105.0207
11.6564 | -105.0412
11.66 | -105.064
11.6637 | -105.0842
11.6674 | -105.0964
11.671 | -105.11
11.6747 | -105.119
11.6784 | -105.1235
11.682 | -105.128
11.6857 | -105.134
11.6893 | -105.1331
11.693 | -105.1302
11.6967 | -105.1257
11.7003 | -105.1193
11.704 | -105.1137
11.7077 | -105.1072
11.7113 | -105.0984
11.715 | -105.0913
11.7187 | -105.0849
11.7223 | -105.0844
11.726 | -105.0841
11.7297 | -105.0829
11.7333 | -105.0821
11.737 | -105.0847
11.7407 | -105.0861
11.7443 | -105.0881
11.748 | -105.0922
11.7517 | -105.0983
11.7553 | -105.1045
11.759 | -105.1163
11.7627 | -105.1292
11.7663 | -105.1397
11.77 | -105.1502
11.7737 | -105.1625
11.7773 | -105.1741
11.781 | -105.1854
11.7847 | -105.1969
11.7883 | -105.2089
11.792 | -105.2213
11.7956 | -105.2324
11.7993 | -105.2444
11.803 | -105.2531
11.8066 | -105.264
11.8103 | -105.2775
11.814 | -105.2864
11.8176 | -105.2952
11.8213 | -105.3023
11.825 | -105.3155
11.8286 | -105.3254
11.8323 | -105.3304
11.836 | -105.3397
11.8396 | -105.3517
11.8433 | -105.363
11.847 | -105.377
11.8506 | -105.3904
11.8543 | -105.4021
11.858 | -105.4123
11.8616 | -105.4224
11.8653 | -105.429
11.869 | -105.4392
11.8726 | -105.4471
11.8763 | -105.4556
11.88 | -105.4666
11.8836 | -105.4731
11.8873 | -105.4767
11.891 | -105.479
11.8946 | -105.4757
11.8983 | -105.4733
11.9019 | -105.4695
11.9056 | -105.4667
11.9093 | -105.4621
11.9129 | -105.4591
11.9166 | -105.4531
11.9203 | -105.4444
11.9239 | -105.4324
11.9276 | -105.4174
11.9313 | -105.4024
11.9349 | -105.3886
11.9386 | -105.3728
11.9423 | -105.3558
11.9459 | -105.3378
11.9496 | -105.3181
11.9533 | -105.299
11.9569 | -105.2796
11.9606 | -105.2583
11.9643 | -105.2362
11.9679 | -105.2169
11.9716 | -105.1952
11.9753 | -105.1746
11.9789 | -105.1557
11.9826 | -105.1372
11.9863 | -105.1244
11.9899 | -105.1087
11.9936 | -105.0973
11.9973 | -105.0852
12.0009 | -105.075
12.0046 | -105.0706
12.0082 | -105.069
12.0119 | -105.0668
12.0156 | -105.0649
12.0192 | -105.0634
12.0229 | -105.0668
12.0266 | -105.0714
12.0302 | -105.0743
12.0339 | -105.0757
12.0376 | -105.0793
12.0412 | -105.0804
12.0449 | -105.0835
12.0486 | -105.0826
12.0522 | -105.0781
12.0559 | -105.0777
12.0596 | -105.0744
12.0632 | -105.0707
12.0669 | -105.0654
12.0706 | -105.0579
12.0742 | -105.0508
12.0779 | -105.0426
12.0816 | -105.0322
12.0852 | -105.0229
12.0889 | -105.0147
12.0926 | -105.0062
12.0962 | -105.0009
12.0999 | -104.9988
12.1036 | -104.996
12.1072 | -104.9943
12.1109 | -104.995
12.1145 | -105.0017
12.1182 | -105.0104
12.1219 | -105.02
12.1255 | -105.0319
12.1292 | -105.0435
12.1329 | -105.0604
12.1365 | -105.081
12.1402 | -105.1002
12.1439 | -105.1196
12.1475 | -105.136
12.1512 | -105.1556
12.1549 | -105.1677
12.1585 | -105.1761
12.1622 | -105.1816
12.1659 | -105.197
12.1695 | -105.207
12.1732 | -105.2126
12.1769 | -105.2181
12.1805 | -105.2214
12.1842 | -105.2219
12.1879 | -105.2226
12.1915 | -105.2226
12.1952 | -105.2185
12.1989 | -105.212
12.2025 | -105.205
12.2062 | -105.1995
12.2099 | -105.193
12.2135 | -105.1816
12.2172 | -105.1698
12.2208 | -105.1542
12.2245 | -105.141
12.2282 | -105.132
12.2318 | -105.1191
12.2355 | -105.1099
12.2392 | -105.1007
12.2428 | -105.0907
12.2465 | -105.0817
12.2502 | -105.0727
12.2538 | -105.0652
12.2575 | -105.0587
12.2612 | -105.0516
12.2648 | -105.0415
12.2685 | -105.0317
12.2722 | -105.0201
12.2758 | -105.0072
12.2795 | -104.9919
12.2832 | -104.9786
12.2868 | -104.9665
12.2905 | -104.9526
12.2942 | -104.9388
12.2978 | -104.9225
12.3015 | -104.9059
12.3052 | -104.8954
12.3088 | -104.8832
12.3125 | -104.8717
12.3162 | -104.8606
12.3198 | -104.8506
12.3235 | -104.8457
12.3271 | -104.841
12.3308 | -104.845
12.3345 | -104.8485
12.3381 | -104.8505
12.3418 | -104.8528
12.3455 | -104.8557
12.3491 | -104.8662
12.3528 | -104.8825
12.3565 | -104.9014
12.3601 | -104.92
12.3638 | -104.9406
12.3675 | -104.9624
12.3711 | -104.9841
12.3748 | -105.0098
12.3785 | -105.0358
12.3821 | -105.0629
12.3858 | -105.0885
12.3895 | -105.115
12.3931 | -105.1397
12.3968 | -105.1588
12.4005 | -105.177
12.4041 | -105.1938
12.4078 | -105.2067
12.4115 | -105.2193
12.4151 | -105.2314
12.4188 | -105.2401
12.4225 | -105.2463
12.4261 | -105.2478
12.4298 | -105.2477
12.4334 | -105.2467
12.4371 | -105.2419
12.4408 | -105.2327
12.4444 | -105.2266
12.4481 | -105.2177
12.4518 | -105.205
12.4554 | -105.1907
12.4591 | -105.1745
12.4628 | -105.1529
12.4664 | -105.1297
12.4701 | -105.104
12.4738 | -105.0811
12.4774 | -105.0583
12.4811 | -105.0342
12.4848 | -105.0096
12.4884 | -104.9865
12.4921 | -104.966
12.4958 | -104.9506
12.4994 | -104.9357
12.5031 | -104.9227
12.5068 | -104.9109
12.5104 | -104.9014
12.5141 | -104.897
12.5178 | -104.8924
12.5214 | -104.8894
12.5251 | -104.8932
12.5288 | -104.8956
12.5324 | -104.9012
12.5361 | -104.905
12.5397 | -104.9131
12.5434 | -104.9206
12.5471 | -104.9318
12.5507 | -104.9441
12.5544 | -104.9562
12.5581 | -104.9664
12.5617 | -104.976
12.5654 | -104.9865
12.5691 | -104.9971
12.5727 | -105.0075
12.5764 | -105.0179
12.5801 | -105.0261
12.5837 | -105.034
12.5874 | -105.0406
12.5911 | -105.0467
12.5947 | -105.0484
12.5984 | -105.0522
12.6021 | -105.0571
12.6057 | -105.0618
12.6094 | -105.0677
12.6131 | -105.0721
12.6167 | -105.0819
12.6204 | -105.0925
12.6241 | -105.0971
12.6277 | -105.1058
12.6314 | -105.1185
12.6351 | -105.1309
12.6387 | -105.1444
12.6424 | -105.1613
12.646 | -105.1771
12.6497 | -105.1904
12.6534 | -105.2045
12.657 | -105.2258
12.6607 | -105.2469
12.6644 | -105.2636
12.668 | -105.2798
12.6717 | -105.2957
12.6754 | -105.3123
12.679 | -105.3302
12.6827 | -105.3428
12.6864 | -105.3545
12.69 | -105.3539
12.6937 | -105.362
12.6974 | -105.3664
12.701 | -105.3719
12.7047 | -105.3749
12.7084 | -105.3714
12.712 | -105.3619
12.7157 | -105.3519
12.7194 | -105.3446
12.723 | -105.3321
12.7267 | -105.3225
12.7304 | -105.3069
12.734 | -105.2898
12.7377 | -105.2723
12.7414 | -105.2532
12.745 | -105.2332
12.7487 | -105.2131
12.7523 | -105.1925
12.756 | -105.1762
12.7597 | -105.1623
12.7633 | -105.1476
12.767 | -105.1309
12.7707 | -105.1147
12.7743 | -105.099
12.778 | -105.0851
12.7817 | -105.0728
12.7853 | -105.0616
12.789 | -105.051
12.7927 | -105.037
12.7963 | -105.0241
12.8 | -105.011
12.8037 | -105.0018
12.8073 | -104.9926
12.811 | -104.981
12.8147 | -104.9697
12.8183 | -104.9593
12.822 | -104.9513
12.8257 | -104.9454
12.8293 | -104.9436
12.833 | -104.9463
12.8367 | -104.9466
12.8403 | -104.9452
12.844 | -104.9462
12.8477 | -104.9514
12.8513 | -104.9599
12.855 | -104.9676
12.8586 | -104.9742
12.8623 | -104.9827
12.866 | -104.9925
12.8696 | -105.0059
12.8733 | -105.0198
12.877 | -105.0311
12.8806 | -105.0419
12.8843 | -105.0541
12.888 | -105.0652
12.8916 | -105.0759
12.8953 | -105.0853
12.899 | -105.0995
12.9026 | -105.1116
12.9063 | -105.1176
12.91 | -105.1238
12.9136 | -105.1303
12.9173 | -105.1388
12.921 | -105.1447
12.9246 | -105.147
12.9283 | -105.1515
12.932 | -105.1575
12.9356 | -105.1605
12.9393 | -105.1651
12.943 | -105.1681
12.9466 | -105.1673
12.9503 | -105.1677
12.954 | -105.1674
12.9576 | -105.1654
12.9613 | -105.1652
12.9649 | -105.1634
12.9686 | -105.1618
12.9723 | -105.1586
12.9759 | -105.1556
12.9796 | -105.1557
12.9833 | -105.1587
12.9869 | -105.1603
12.9906 | -105.1575
12.9943 | -105.1591
12.9979 | -105.1661
13.0016 | -105.173
13.0053 | -105.1794
13.0089 | -105.1852
13.0126 | -105.19
13.0163 | -105.1951
13.0199 | -105.1998
13.0236 | -105.2035
13.0273 | -105.209
13.0309 | -105.2163
13.0346 | -105.2231
13.0383 | -105.2258
13.0419 | -105.2318
13.0456 | -105.2408
13.0493 | -105.2504
13.0529 | -105.2576
13.0566 | -105.2629
13.0603 | -105.2712
13.0639 | -105.2822
13.0676 | -105.2935
13.0712 | -105.3021
13.0749 | -105.3112
13.0786 | -105.3235
13.0822 | -105.3379
13.0859 | -105.349
13.0896 | -105.3618
13.0932 | -105.3819
13.0969 | -105.4003
13.1006 | -105.4106
13.1042 | -105.4183
13.1079 | -105.4271
13.1116 | -105.4359
13.1152 | -105.4417
13.1189 | -105.4427
13.1226 | -105.4391
13.1262 | -105.4301
13.1299 | -105.4175
13.1336 | -105.4027
13.1372 | -105.3874
13.1409 | -105.3723
13.1446 | -105.3553
13.1482 | -105.3315
13.1519 | -105.3081
13.1556 | -105.2846
13.1592 | -105.2623
13.1629 | -105.2396
13.1666 | -105.2165
13.1702 | -105.1892
13.1739 | -105.1632
13.1775 | -105.1459
13.1812 | -105.1284
13.1849 | -105.1088
13.1885 | -105.0929
13.1922 | -105.0794
13.1959 | -105.0693
13.1995 | -105.0644
13.2032 | -105.0623
13.2069 | -105.0607
13.2105 | -105.0631
13.2142 | -105.0696
13.2179 | -105.0905
13.2215 | -105.0988
13.2252 | -105.108
13.2289 | -105.1196
13.2325 | -105.1301
13.2362 | -105.141
13.2399 | -105.1554
13.2435 | -105.1714
13.2472 | -105.1861
13.2509 | -105.1995
13.2545 | -105.2125
13.2582 | -105.2237
13.2619 | -105.2338
13.2655 | -105.2457
13.2692 | -105.2545
13.2729 | -105.2614
13.2765 | -105.2649
13.2802 | -105.2641
13.2838 | -105.2622
13.2875 | -105.259
13.2912 | -105.2555
13.2948 | -105.2556
13.2985 | -105.2542
13.3022 | -105.25
13.3058 | -105.247
13.3095 | -105.2436
13.3132 | -105.2419
13.3168 | -105.2384
13.3205 | -105.2337
13.3242 | -105.2321
13.3278 | -105.2286
13.3315 | -105.2241
13.3352 | -105.2206
13.3388 | -105.2207
13.3425 | -105.2195
13.3462 | -105.2179
13.3498 | -105.2173
13.3535 | -105.2169
13.3572 | -105.218
13.3608 | -105.2219
13.3645 | -105.2293
13.3682 | -105.2322
13.3718 | -105.2359
13.3755 | -105.242
13.3792 | -105.2501
13.3828 | -105.2611
13.3865 | -105.2714
13.3901 | -105.28
13.3938 | -105.2923
13.3975 | -105.3044
13.4011 | -105.3122
13.4048 | -105.3174
13.4085 | -105.3225
13.4121 | -105.3312
13.4158 | -105.3366
13.4195 | -105.3431
13.4231 | -105.3474
13.4268 | -105.3507
13.4305 | -105.3506
13.4341 | -105.3482
13.4378 | -105.3466
13.4415 | -105.3434
13.4451 | -105.3411
13.4488 | -105.3376
13.4525 | -105.3318
13.4561 | -105.3283
13.4598 | -105.322
13.4635 | -105.317
13.4671 | -105.3142
13.4708 | -105.3121
13.4745 | -105.3122
13.4781 | -105.3133
13.4818 | -105.3164
13.4855 | -105.323
13.4891 | -105.3256
13.4928 | -105.3282
13.4964 | -105.3314
13.5001 | -105.3345
13.5038 | -105.3387
13.5074 | -105.3454
13.5111 | -105.3552
13.5148 | -105.3649
13.5184 | -105.3724
13.5221 | -105.3804
13.5258 | -105.3881
13.5294 | -105.3952
13.5331 | -105.4003
13.5368 | -105.4066
13.5404 | -105.416
13.5441 | -105.4195
13.5478 | -105.4215
13.5514 | -105.4262
13.5551 | -105.4275
13.5588 | -105.427
13.5624 | -105.4306
13.5661 | -105.4304
13.5698 | -105.4272
13.5734 | -105.4256
13.5771 | -105.4228
13.5808 | -105.4203
13.5844 | -105.4161
13.5881 | -105.4096
13.5918 | -105.4057
13.5954 | -105.4011
13.5991 | -105.3942
13.6027 | -105.3862
13.6064 | -105.3735
13.6101 | -105.3638
13.6137 | -105.3532
13.6174 | -105.347
13.6211 | -105.3393
13.6247 | -105.3325
13.6284 | -105.3261
13.6321 | -105.3191
13.6357 | -105.3161
13.6394 | -105.3159
13.6431 | -105.3157
13.6467 | -105.3197
13.6504 | -105.3234
13.6541 | -105.3269
13.6577 | -105.3291
13.6614 | -105.331
13.6651 | -105.3353
13.6687 | -105.3403
13.6724 | -105.3434
13.6761 | -105.3458
13.6797 | -105.3482
13.6834 | -105.3522
13.6871 | -105.3567
13.6907 | -105.3614
13.6944 | -105.3657
13.6981 | -105.3693
13.7017 | -105.3691
13.7054 | -105.3671
13.709 | -105.3655
13.7127 | -105.363
13.7164 | -105.3612
13.72 | -105.3588
13.7237 | -105.358
13.7274 | -105.3568
13.731 | -105.3543
13.7347 | -105.3495
13.7384 | -105.3465
13.742 | -105.3413
13.7457 | -105.3345
13.7494 | -105.329
13.753 | -105.3252
13.7567 | -105.3196
13.7604 | -105.311
13.764 | -105.3038
13.7677 | -105.2957
13.7714 | -105.2833
13.775 | -105.2701
13.7787 | -105.2598
13.7824 | -105.2513
13.786 | -105.2392
13.7897 | -105.2294
13.7934 | -105.2195
13.797 | -105.2092
13.8007 | -105.1986
13.8044 | -105.1887
13.808 | -105.172
13.8117 | -105.1499
13.8153 | -105.1234
13.819 | -105.0954
13.8227 | -105.0685
13.8263 | -105.0412
13.83 | -105.0151
13.8337 | -104.9855
13.8373 | -104.954
13.841 | -104.9248
13.8447 | -104.8954
13.8483 | -104.8626
13.852 | -104.8328
13.8557 | -104.8049
13.8593 | -104.7757
13.863 | -104.7492
13.8667 | -104.723
13.8703 | -104.7002
13.874 | -104.6804
13.8777 | -104.6609
13.8813 | -104.6414
13.885 | -104.6259
13.8887 | -104.6169
13.8923 | -104.6083
13.896 | -104.6031
13.8997 | -104.6041
13.9033 | -104.609
13.907 | -104.6159
13.9107 | -104.6281
13.9143 | -104.6411
13.918 | -104.6586
13.9216 | -104.6799
13.9253 | -104.7044
13.929 | -104.734
13.9326 | -104.7648
13.9363 | -104.7989
13.94 | -104.8354
13.9436 | -104.8723
13.9473 | -104.9094
13.951 | -104.9476
13.9546 | -104.9877
13.9583 | -105.0262
13.962 | -105.0631
13.9656 | -105.1005
13.9693 | -105.1373
13.973 | -105.1734
13.9766 | -105.2083
13.9803 | -105.2358
13.984 | -105.263
13.9876 | -105.2924
13.9913 | -105.319
13.995 | -105.3442
13.9986 | -105.3666
14.0023 | -105.3848
14.006 | -105.403
14.0096 | -105.4247
14.0133 | -105.4487
14.017 | -105.4703
14.0206 | -105.489
14.0243 | -105.502
14.028 | -105.51
14.0316 | -105.5176
14.0353 | -105.5269
14.0389 | -105.5318
14.0426 | -105.5344
14.0463 | -105.538
14.0499 | -105.5457
14.0536 | -105.5515
14.0573 | -105.5572
14.0609 | -105.5629
14.0646 | -105.5695
14.0683 | -105.5771
14.0719 | -105.5856
14.0756 | -105.5927
14.0793 | -105.598
14.0829 | -105.6068
14.0866 | -105.6146
14.0903 | -105.6178
14.0939 | -105.6278
14.0976 | -105.6377
14.1013 | -105.644
14.1049 | -105.6491
14.1086 | -105.6547
14.1123 | -105.657
14.1159 | -105.662
14.1196 | -105.6632
14.1233 | -105.6627
14.1269 | -105.6611
14.1306 | -105.6558
14.1342 | -105.6516
14.1379 | -105.6487
14.1416 | -105.6454
14.1452 | -105.6393
14.1489 | -105.6345
14.1526 | -105.6306
14.1562 | -105.6252
14.1599 | -105.6206
14.1636 | -105.6167
14.1672 | -105.6149
14.1709 | -105.6148
14.1746 | -105.6153
14.1782 | -105.6153
14.1819 | -105.6178
14.1856 | -105.6179
14.1892 | -105.6205
14.1929 | -105.6225
14.1966 | -105.622
14.2002 | -105.6229
14.2039 | -105.6268
14.2076 | -105.6307
14.2112 | -105.6305
14.2149 | -105.6287
14.2186 | -105.6266
14.2222 | -105.6268
14.2259 | -105.625
14.2296 | -105.6215
14.2332 | -105.613
14.2369 | -105.6028
14.2406 | -105.5935
14.2442 | -105.5834
14.2479 | -105.5678
14.2515 | -105.5469
14.2552 | -105.5246
14.2589 | -105.5013
14.2625 | -105.4791
14.2662 | -105.4546
14.2699 | -105.4284
14.2735 | -105.4018
14.2772 | -105.3807
14.2809 | -105.3584
14.2845 | -105.3399
14.2882 | -105.3201
14.2919 | -105.303
14.2955 | -105.2889
14.2992 | -105.2744
14.3029 | -105.2626
14.3065 | -105.2527
14.3102 | -105.2462
14.3139 | -105.241
14.3175 | -105.2384
14.3212 | -105.2354
14.3249 | -105.2353
14.3285 | -105.2378
14.3322 | -105.2455
14.3359 | -105.2516
14.3395 | -105.2567
14.3432 | -105.2642
14.3468 | -105.2728
14.3505 | -105.281
14.3542 | -105.2871
14.3578 | -105.293
14.3615 | -105.2982
14.3652 | -105.305
14.3688 | -105.3132
14.3725 | -105.3193
14.3762 | -105.3208
14.3798 | -105.3181
14.3835 | -105.3177
14.3872 | -105.3157
14.3908 | -105.3114
14.3945 | -105.3026
14.3982 | -105.2973
14.4018 | -105.2935
14.4055 | -105.2848
14.4092 | -105.274
14.4128 | -105.2642
14.4165 | -105.2574
14.4202 | -105.2503
14.4238 | -105.243
14.4275 | -105.2322
14.4312 | -105.2252
14.4348 | -105.2205
14.4385 | -105.2173
14.4422 | -105.2116
14.4458 | -105.2025
14.4495 | -105.1948
14.4532 | -105.1873
14.4568 | -105.1806
14.4605 | -105.1742
14.4641 | -105.1687
14.4678 | -105.1627
14.4715 | -105.1555
14.4751 | -105.1513
14.4788 | -105.1481
14.4825 | -105.1456
14.4861 | -105.1408
14.4898 | -105.1341
14.4935 | -105.1284
14.4971 | -105.1262
14.5008 | -105.1249
14.5045 | -105.1252
14.5081 | -105.1251
14.5118 | -105.1272
14.5155 | -105.1312
14.5191 | -105.1353
14.5228 | -105.1397
14.5265 | -105.145
14.5301 | -105.1529
14.5338 | -105.1616
14.5375 | -105.1669
14.5411 | -105.169
14.5448 | -105.176
14.5485 | -105.185
14.5521 | -105.1938
14.5558 | -105.2009
14.5595 | -105.2095
14.5631 | -105.217
14.5668 | -105.2246
14.5704 | -105.2335
14.5741 | -105.239
14.5778 | -105.2428
14.5814 | -105.2485
14.5851 | -105.2546
14.5888 | -105.264
14.5924 | -105.2724
14.5961 | -105.282
14.5998 | -105.2899
14.6034 | -105.2986
14.6071 | -105.3034
14.6108 | -105.3089
14.6144 | -105.318
14.6181 | -105.3291
14.6218 | -105.3374
14.6254 | -105.3432
14.6291 | -105.3508
14.6328 | -105.3563
14.6364 | -105.3608
14.6401 | -105.3629
14.6438 | -105.3663
14.6474 | -105.369
14.6511 | -105.3724
14.6548 | -105.3767
14.6584 | -105.378
14.6621 | -105.3783
14.6657 | -105.3776
14.6694 | -105.3744
14.6731 | -105.3709
14.6767 | -105.3671
14.6804 | -105.3607
14.6841 | -105.3489
14.6877 | -105.3393
14.6914 | -105.3304
14.6951 | -105.322
14.6987 | -105.3184
14.7024 | -105.3097
14.7061 | -105.2977
14.7097 | -105.2855
14.7134 | -105.2752
14.7171 | -105.2662
14.7207 | -105.2581
14.7244 | -105.2517
14.7281 | -105.2446
14.7317 | -105.2371
14.7354 | -105.2294
14.7391 | -105.2212
14.7427 | -105.2105
14.7464 | -105.2048
14.7501 | -105.1994
14.7537 | -105.1898
14.7574 | -105.18
14.7611 | -105.1693
14.7647 | -105.1622
14.7684 | -105.1564
14.7721 | -105.1538
14.7757 | -105.1521
14.7794 | -105.1489
14.783 | -105.146
14.7867 | -105.1436
14.7904 | -105.1398
14.794 | -105.1401
14.7977 | -105.1432
14.8014 | -105.1452
14.805 | -105.1566
14.8087 | -105.1618
14.8124 | -105.1708
14.816 | -105.181
14.8197 | -105.1926
14.8234 | -105.2055
14.827 | -105.2166
14.8307 | -105.2268
14.8344 | -105.2388
14.838 | -105.2544
14.8417 | -105.2692
14.8454 | -105.2882
14.849 | -105.3081
14.8527 | -105.3246
14.8564 | -105.3372
14.86 | -105.3537
14.8637 | -105.3709
14.8674 | -105.3871
14.871 | -105.4026
14.8747 | -105.4183
14.8784 | -105.4306
14.882 | -105.4403
14.8857 | -105.4467
14.8893 | -105.452
14.893 | -105.4573
14.8967 | -105.4633
14.9003 | -105.4663
14.904 | -105.4652
14.9077 | -105.4635
14.9113 | -105.458
14.915 | -105.4506
14.9187 | -105.4407
14.9223 | -105.4287
14.926 | -105.4155
14.9297 | -105.3979
14.9333 | -105.38
14.937 | -105.3615
14.9407 | -105.342
14.9443 | -105.3218
14.948 | -105.3008
14.9517 | -105.2763
14.9553 | -105.2551
14.959 | -105.2319
14.9627 | -105.2092
14.9663 | -105.1871
14.97 | -105.1645
14.9737 | -105.1421
14.9773 | -105.12
14.981 | -105.0979
14.9847 | -105.0764
14.9883 | -105.0563
14.992 | -105.0406
14.9956 | -105.0291
14.9993 | -105.0192
15.003 | -105.0101
15.0066 | -105.0047
15.0103 | -104.9983
15.014 | -104.9958
15.0176 | -104.9969
15.0213 | -105.002
15.025 | -105.0087
15.0286 | -105.0162
15.0323 | -105.026
15.036 | -105.0393
15.0396 | -105.054
15.0433 | -105.0702
15.047 | -105.0838
15.0506 | -105.0985
15.0543 | -105.1111
15.058 | -105.1259
15.0616 | -105.1391
15.0653 | -105.1519
15.069 | -105.1664
15.0726 | -105.18
15.0763 | -105.1872
15.08 | -105.1953
15.0836 | -105.2013
15.0873 | -105.2074
15.091 | -105.2114
15.0946 | -105.2157
15.0983 | -105.2177
15.1019 | -105.2201
15.1056 | -105.2218
15.1093 | -105.221
15.1129 | -105.2164
15.1166 | -105.213
15.1203 | -105.2128
15.1239 | -105.2135
15.1276 | -105.2142
15.1313 | -105.2167
15.1349 | -105.2194
15.1386 | -105.2232
15.1423 | -105.2273
15.1459 | -105.2345
15.1496 | -105.2391
15.1533 | -105.2445
15.1569 | -105.2532
15.1606 | -105.2638
15.1643 | -105.2739
15.1679 | -105.2889
15.1716 | -105.3034
15.1753 | -105.3183
15.1789 | -105.3312
15.1826 | -105.3459
15.1863 | -105.3618
15.1899 | -105.3776
15.1936 | -105.3919
15.1973 | -105.4095
15.2009 | -105.427
15.2046 | -105.4463
15.2082 | -105.4643
15.2119 | -105.4811
15.2156 | -105.4982
15.2192 | -105.5153
15.2229 | -105.5278
15.2266 | -105.5415
15.2302 | -105.5558
15.2339 | -105.5667
15.2376 | -105.5741
15.2412 | -105.5824
15.2449 | -105.5918
15.2486 | -105.5992
15.2522 | -105.6061
15.2559 | -105.6135
15.2596 | -105.6218
15.2632 | -105.6297
15.2669 | -105.6339
15.2706 | -105.6407
15.2742 | -105.6449
15.2779 | -105.6484
15.2816 | -105.6533
15.2852 | -105.6587
15.2889 | -105.6662
15.2926 | -105.6742
15.2962 | -105.6781
15.2999 | -105.6852
15.3036 | -105.6869
15.3072 | -105.6868
15.3109 | -105.6899
15.3145 | -105.6951
15.3182 | -105.6973
15.3219 | -105.6974
15.3255 | -105.6987
15.3292 | -105.7003
15.3329 | -105.6969
15.3365 | -105.6948
15.3402 | -105.6921
15.3439 | -105.6865
15.3475 | -105.6825
15.3512 | -105.68
15.3549 | -105.6791
15.3585 | -105.6744
15.3622 | -105.6708
15.3659 | -105.6699
15.3695 | -105.6718
15.3732 | -105.6723
15.3769 | -105.6716
15.3805 | -105.6707
15.3842 | -105.6719
15.3879 | -105.6708
15.3915 | -105.6686
15.3952 | -105.6668
15.3989 | -105.6676
15.4025 | -105.6684
15.4062 | -105.6681
15.4099 | -105.6676
15.4135 | -105.6689
15.4172 | -105.6679
15.4208 | -105.6656
15.4245 | -105.6602
15.4282 | -105.6553
15.4318 | -105.65
15.4355 | -105.6448
15.4392 | -105.6414
15.4428 | -105.6361
15.4465 | -105.63
15.4502 | -105.6214
15.4538 | -105.6125
15.4575 | -105.6041
15.4612 | -105.5909
15.4648 | -105.5778
15.4685 | -105.5653
15.4722 | -105.5504
15.4758 | -105.5354
15.4795 | -105.5212
15.4832 | -105.5093
15.4868 | -105.4947
15.4905 | -105.4746
15.4942 | -105.4529
15.4978 | -105.4321
15.5015 | -105.4121
15.5052 | -105.3906
15.5088 | -105.3687
15.5125 | -105.3486
15.5162 | -105.3294
15.5198 | -105.3093
15.5235 | -105.291
15.5271 | -105.2705
15.5308 | -105.2455
15.5345 | -105.2209
15.5381 | -105.1963
15.5418 | -105.1721
15.5455 | -105.1485
15.5491 | -105.1262
15.5528 | -105.1024
15.5565 | -105.0797
15.5601 | -105.0579
15.5638 | -105.0366
15.5675 | -105.018
15.5711 | -105.0019
15.5748 | -104.9846
15.5785 | -104.9684
15.5821 | -104.9578
15.5858 | -104.9479
15.5895 | -104.9401
15.5931 | -104.9347
15.5968 | -104.9337
15.6005 | -104.9356
15.6041 | -104.9402
15.6078 | -104.9465
15.6115 | -104.9565
15.6151 | -104.969
15.6188 | -104.9844
15.6225 | -105.0022
15.6261 | -105.0223
15.6298 | -105.0447
15.6334 | -105.0661
15.6371 | -105.0887
15.6408 | -105.1129
15.6444 | -105.1389
15.6481 | -105.1673
15.6518 | -105.1962
15.6554 | -105.2247
15.6591 | -105.2484
15.6628 | -105.2703
15.6664 | -105.293
15.6701 | -105.3143
15.6738 | -105.3346
15.6774 | -105.3523
15.6811 | -105.3677
15.6848 | -105.3822
15.6884 | -105.3935
15.6921 | -105.4021
15.6958 | -105.4098
15.6994 | -105.4149
15.7031 | -105.415
15.7068 | -105.4124
15.7104 | -105.4127
15.7141 | -105.4122
15.7178 | -105.4076
15.7214 | -105.402
15.7251 | -105.3942
15.7288 | -105.3835
15.7324 | -105.3717
15.7361 | -105.3585
15.7397 | -105.3463
15.7434 | -105.3345
15.7471 | -105.3227
15.7507 | -105.3155
15.7544 | -105.3071
15.7581 | -105.2981
15.7617 | -105.2921
15.7654 | -105.2861
15.7691 | -105.2819
15.7727 | -105.2784
15.7764 | -105.2768
15.7801 | -105.2787
15.7837 | -105.2796
15.7874 | -105.282
15.7911 | -105.2853
15.7947 | -105.289
15.7984 | -105.294
15.8021 | -105.3005
15.8057 | -105.3038
15.8094 | -105.3081
15.8131 | -105.3115
15.8167 | -105.3133
15.8204 | -105.3163
15.8241 | -105.3204
15.8277 | -105.3228
15.8314 | -105.3254
15.8351 | -105.327
15.8387 | -105.3275
15.8424 | -105.3267
15.846 | -105.3254
15.8497 | -105.3207
15.8534 | -105.3146
15.857 | -105.3116
15.8607 | -105.3059
15.8644 | -105.2968
15.868 | -105.2904
15.8717 | -105.2858
15.8754 | -105.2791
15.879 | -105.27
15.8827 | -105.2607
15.8864 | -105.256
15.89 | -105.2495
15.8937 | -105.242
15.8974 | -105.2349
15.901 | -105.2289
15.9047 | -105.2226
15.9084 | -105.2149
15.912 | -105.2029
15.9157 | -105.1923
15.9194 | -105.1816
15.923 | -105.1726
15.9267 | -105.1645
15.9304 | -105.1597
15.934 | -105.1541
15.9377 | -105.1463
15.9414 | -105.1399
15.945 | -105.136
15.9487 | -105.1322
15.9523 | -105.127
15.956 | -105.1257
15.9597 | -105.1259
15.9633 | -105.1274
15.967 | -105.1283
15.9707 | -105.1299
15.9743 | -105.1323
15.978 | -105.1298
15.9817 | -105.1328
15.9853 | -105.1353
15.989 | -105.1369
15.9927 | -105.14
15.9963 | -105.1387
16 | -105.1404
16.0037 | -105.1404
16.0073 | -105.1384
16.011 | -105.1375
16.0147 | -105.1392
16.0183 | -105.1414
16.022 | -105.1415
16.0257 | -105.1406
16.0293 | -105.1365
16.033 | -105.1346
16.0367 | -105.1335
16.0403 | -105.1296
16.044 | -105.125
16.0477 | -105.1179
16.0513 | -105.1081
16.055 | -105.1006
16.0586 | -105.0916
16.0623 | -105.0797
16.066 | -105.0691
16.0696 | -105.0577
16.0733 | -105.0464
16.077 | -105.0355
16.0806 | -105.0241
16.0843 | -105.0097
16.088 | -104.9957
16.0916 | -104.9799
16.0953 | -104.9663
16.099 | -104.9522
16.1026 | -104.9351
16.1063 | -104.9204
16.11 | -104.9102
16.1136 | -104.902
16.1173 | -104.8935
16.121 | -104.8863
16.1246 | -104.8813
16.1283 | -104.8764
16.132 | -104.8722
16.1356 | -104.8703
16.1393 | -104.872
16.143 | -104.8792
16.1466 | -104.8857
16.1503 | -104.8923
16.154 | -104.9003
16.1576 | -104.9098
16.1613 | -104.9197
16.1649 | -104.9323
16.1686 | -104.9473
16.1723 | -104.9633
16.1759 | -104.9818
16.1796 | -104.9999
16.1833 | -105.0196
16.1869 | -105.0383
16.1906 | -105.056
16.1943 | -105.0719
16.1979 | -105.0897
16.2016 | -105.1079
16.2053 | -105.125
16.2089 | -105.1445
16.2126 | -105.1596
16.2163 | -105.1733
16.2199 | -105.1862
16.2236 | -105.1966
16.2273 | -105.2075
16.2309 | -105.2147
16.2346 | -105.222
16.2383 | -105.2273
16.2419 | -105.2306
16.2456 | -105.2319
16.2493 | -105.2309
16.2529 | -105.2311
16.2566 | -105.233
16.2603 | -105.2339
16.2639 | -105.2345
16.2676 | -105.2342
16.2712 | -105.2355
16.2749 | -105.2361
16.2786 | -105.2349
16.2822 | -105.2359
16.2859 | -105.2361
16.2896 | -105.2378
16.2932 | -105.2406
16.2969 | -105.2403
16.3006 | -105.2407
16.3042 | -105.2416
16.3079 | -105.2431
16.3116 | -105.2449
16.3152 | -105.2496
16.3189 | -105.2549
16.3226 | -105.2623
16.3262 | -105.2684
16.3299 | -105.2756
16.3336 | -105.2844
16.3372 | -105.2892
16.3409 | -105.2927
16.3446 | -105.2963
16.3482 | -105.3007
16.3519 | -105.3071
16.3556 | -105.3124
16.3592 | -105.3186
16.3629 | -105.324
16.3666 | -105.3295
16.3702 | -105.3348
16.3739 | -105.3392
16.3775 | -105.344
16.3812 | -105.3478
16.3849 | -105.3514
16.3885 | -105.3562
16.3922 | -105.3573
16.3959 | -105.3613
16.3995 | -105.3628
16.4032 | -105.3632
16.4069 | -105.3633
16.4105 | -105.3605
16.4142 | -105.3549
16.4179 | -105.3478
16.4215 | -105.342
16.4252 | -105.3359
16.4289 | -105.3317
16.4325 | -105.3264
16.4362 | -105.3198
16.4399 | -105.3146
16.4435 | -105.3086
16.4472 | -105.3011
16.4509 | -105.2956
16.4545 | -105.2886
16.4582 | -105.2825
16.4619 | -105.2791
16.4655 | -105.2739
16.4692 | -105.269
16.4729 | -105.2668
16.4765 | -105.2649
16.4802 | -105.2626
16.4838 | -105.2614
16.4875 | -105.2607
16.4912 | -105.2593
16.4948 | -105.2605
16.4985 | -105.2607
16.5022 | -105.2621
16.5058 | -105.2649
16.5095 | -105.2687
16.5132 | -105.2753
16.5168 | -105.2807
16.5205 | -105.2883
16.5242 | -105.2986
16.5278 | -105.3106
16.5315 | -105.323
16.5352 | -105.3352
16.5388 | -105.3481
16.5425 | -105.3636
16.5462 | -105.3813
16.5498 | -105.3963
16.5535 | -105.413
16.5572 | -105.43
16.5608 | -105.4484
16.5645 | -105.468
16.5682 | -105.4896
16.5718 | -105.5134
16.5755 | -105.5375
16.5792 | -105.5608
16.5828 | -105.5843
16.5865 | -105.6091
16.5901 | -105.634
16.5938 | -105.6582
16.5975 | -105.6803
16.6011 | -105.7018
16.6048 | -105.7213
16.6085 | -105.7372
16.6121 | -105.7514
16.6158 | -105.7667
16.6195 | -105.7805
16.6231 | -105.7924
16.6268 | -105.8027
16.6305 | -105.8118
16.6341 | -105.8183
16.6378 | -105.8272
16.6415 | -105.8349
16.6451 | -105.8371
16.6488 | -105.8363
16.6525 | -105.8382
16.6561 | -105.839
16.6598 | -105.8358
16.6635 | -105.8317
16.6671 | -105.8268
16.6708 | -105.8206
16.6745 | -105.8129
16.6781 | -105.8008
16.6818 | -105.7869
16.6855 | -105.7766
16.6891 | -105.7661
16.6928 | -105.756
16.6964 | -105.7442
16.7001 | -105.7315
16.7038 | -105.7229
16.7074 | -105.7124
16.7111 | -105.7011
16.7148 | -105.6914
16.7184 | -105.6811
16.7221 | -105.6713
16.7258 | -105.6622
16.7294 | -105.6557
16.7331 | -105.6482
16.7368 | -105.6432
16.7404 | -105.6372
16.7441 | -105.6311
16.7478 | -105.6272
16.7514 | -105.6236
16.7551 | -105.6216
16.7588 | -105.6188
16.7624 | -105.6161
16.7661 | -105.6154
16.7698 | -105.616
16.7734 | -105.6165
16.7771 | -105.6174
16.7808 | -105.6185
16.7844 | -105.6169
16.7881 | -105.6173
16.7918 | -105.6166
16.7954 | -105.6165
16.7991 | -105.6156
16.8027 | -105.6157
16.8064 | -105.6177
16.8101 | -105.6184
16.8137 | -105.6153
16.8174 | -105.6129
16.8211 | -105.6086
16.8247 | -105.6036
16.8284 | -105.5988
16.8321 | -105.5949
16.8357 | -105.5915
16.8394 | -105.5875
16.8431 | -105.5819
16.8467 | -105.5759
16.8504 | -105.5682
16.8541 | -105.5592
16.8577 | -105.548
16.8614 | -105.5378
16.8651 | -105.5281
16.8687 | -105.5188
16.8724 | -105.5103
16.8761 | -105.4973
16.8797 | -105.485
16.8834 | -105.4742
16.8871 | -105.4668
16.8907 | -105.4558
16.8944 | -105.4459
16.8981 | -105.4358
16.9017 | -105.4259
16.9054 | -105.4139
16.909 | -105.4005
16.9127 | -105.3896
16.9164 | -105.3831
16.92 | -105.3773
16.9237 | -105.3716
16.9274 | -105.3617
16.931 | -105.3531
16.9347 | -105.3446
16.9384 | -105.3371
16.942 | -105.3299
16.9457 | -105.3223
16.9494 | -105.319
16.953 | -105.3174
16.9567 | -105.3167
16.9604 | -105.3169
16.964 | -105.3174
16.9677 | -105.318
16.9714 | -105.3212
16.975 | -105.3246
16.9787 | -105.3265
16.9824 | -105.331
16.986 | -105.3369
16.9897 | -105.344
16.9934 | -105.3521
16.997 | -105.3599
17.0007 | -105.368
17.0044 | -105.3754
17.008 | -105.3843
17.0117 | -105.3951
17.0153 | -105.4062
17.019 | -105.4151
17.0227 | -105.4234
17.0263 | -105.4353
17.03 | -105.447
17.0337 | -105.4573
17.0373 | -105.4659
17.041 | -105.4753
17.0447 | -105.4845
17.0483 | -105.492
17.052 | -105.501
17.0557 | -105.5113
17.0593 | -105.5208
17.063 | -105.5292
17.0667 | -105.536
17.0703 | -105.5425
17.074 | -105.5493
17.0777 | -105.5574
17.0813 | -105.5649
17.085 | -105.5704
17.0887 | -105.576
17.0923 | -105.5803
17.096 | -105.5844
17.0997 | -105.5872
17.1033 | -105.5884
17.107 | -105.5893
17.1107 | -105.5903
17.1143 | -105.5891
17.118 | -105.5898
17.1216 | -105.5886
17.1253 | -105.5864
17.129 | -105.5814
17.1326 | -105.5776
17.1363 | -105.5736
17.14 | -105.5691
17.1436 | -105.5624
17.1473 | -105.5548
17.151 | -105.548
17.1546 | -105.5401
17.1583 | -105.5311
17.162 | -105.5233
17.1656 | -105.5179
17.1693 | -105.5105
17.173 | -105.5016
17.1766 | -105.4923
17.1803 | -105.4846
17.184 | -105.4773
17.1876 | -105.4696
17.1913 | -105.4609
17.195 | -105.4515
17.1986 | -105.4408
17.2023 | -105.4306
17.206 | -105.4215
17.2096 | -105.4146
17.2133 | -105.4071
17.217 | -105.3972
17.2206 | -105.3897
17.2243 | -105.3851
17.2279 | -105.3806
17.2316 | -105.3763
17.2353 | -105.369
17.2389 | -105.3615
17.2426 | -105.3554
17.2463 | -105.35
17.2499 | -105.3434
17.2536 | -105.3361
17.2573 | -105.3283
17.2609 | -105.3181
17.2646 | -105.3099
17.2683 | -105.3003
17.2719 | -105.287
17.2756 | -105.2729
17.2793 | -105.2582
17.2829 | -105.2431
17.2866 | -105.2259
17.2903 | -105.2082
17.2939 | -105.1885
17.2976 | -105.1697
17.3013 | -105.1516
17.3049 | -105.1319
17.3086 | -105.1119
17.3123 | -105.0922
17.3159 | -105.073
17.3196 | -105.0512
17.3233 | -105.0259
17.3269 | -105.0051
17.3306 | -104.9878
17.3343 | -104.9711
17.3379 | -104.955
17.3416 | -104.9375
17.3452 | -104.9206
17.3489 | -104.9024
17.3526 | -104.8875
17.3562 | -104.8738
17.3599 | -104.863
17.3636 | -104.8529
17.3672 | -104.8448
17.3709 | -104.8373
17.3746 | -104.8301
17.3782 | -104.8249
17.3819 | -104.8184
17.3856 | -104.815
17.3892 | -104.8142
17.3929 | -104.8173
17.3966 | -104.8227
17.4002 | -104.8278
17.4039 | -104.8327
17.4076 | -104.8379
17.4112 | -104.8436
17.4149 | -104.8507
17.4186 | -104.8628
17.4222 | -104.8739
17.4259 | -104.8863
17.4296 | -104.8997
17.4332 | -104.9117
17.4369 | -104.9278
17.4406 | -104.9443
17.4442 | -104.9607
17.4479 | -104.9832
17.4515 | -105.0008
17.4552 | -105.0217
17.4589 | -105.0411
17.4625 | -105.0635
17.4662 | -105.0864
17.4699 | -105.1065
17.4735 | -105.1246
17.4772 | -105.1418
17.4809 | -105.1603
17.4845 | -105.1769
17.4882 | -105.1915
17.4919 | -105.2072
17.4955 | -105.2225
17.4992 | -105.2383
17.5029 | -105.2518
17.5065 | -105.2622
17.5102 | -105.2737
17.5139 | -105.2833
17.5175 | -105.2905
17.5212 | -105.2983
17.5249 | -105.3028
17.5285 | -105.3071
17.5322 | -105.309
17.5359 | -105.3104
17.5395 | -105.3089
17.5432 | -105.3053
17.5468 | -105.3003
17.5505 | -105.2934
17.5542 | -105.2863
17.5578 | -105.2794
17.5615 | -105.2701
17.5652 | -105.2602
17.5688 | -105.2525
17.5725 | -105.2463
17.5762 | -105.235
17.5798 | -105.2235
17.5835 | -105.2131
17.5872 | -105.2019
17.5908 | -105.1935
17.5945 | -105.1891
17.5982 | -105.187
17.6018 | -105.181
17.6055 | -105.1752
17.6092 | -105.1716
17.6128 | -105.1682
17.6165 | -105.1664
17.6202 | -105.1684
17.6238 | -105.1723
17.6275 | -105.1759
17.6312 | -105.1825
17.6348 | -105.1891
17.6385 | -105.1956
17.6422 | -105.2015
17.6458 | -105.2078
17.6495 | -105.2157
17.6532 | -105.2268
17.6568 | -105.2373
17.6605 | -105.2472
17.6641 | -105.2588
17.6678 | -105.2675
17.6715 | -105.2781
17.6751 | -105.2884
17.6788 | -105.2999
17.6825 | -105.315
17.6861 | -105.3246
17.6898 | -105.3341
17.6935 | -105.3461
17.6971 | -105.3565
17.7008 | -105.3663
17.7045 | -105.3778
17.7081 | -105.3883
17.7118 | -105.3997
17.7155 | -105.4085
17.7191 | -105.4176
17.7228 | -105.4272
17.7265 | -105.4376
17.7301 | -105.4475
17.7338 | -105.4549
17.7375 | -105.4646
17.7411 | -105.4747
17.7448 | -105.4788
17.7485 | -105.482
17.7521 | -105.4845
17.7558 | -105.4861
17.7595 | -105.4882
17.7631 | -105.4898
17.7668 | -105.4889
17.7704 | -105.4891
17.7741 | -105.4879
17.7778 | -105.4862
17.7814 | -105.482
17.7851 | -105.4775
17.7888 | -105.4706
17.7924 | -105.4626
17.7961 | -105.4566
17.7998 | -105.4502
17.8034 | -105.4448
17.8071 | -105.4385
17.8108 | -105.4315
17.8144 | -105.4246
17.8181 | -105.4196
17.8218 | -105.4137
17.8254 | -105.4096
17.8291 | -105.4062
17.8328 | -105.4023
17.8364 | -105.3985
17.8401 | -105.3976
17.8438 | -105.3954
17.8474 | -105.3935
17.8511 | -105.392
17.8548 | -105.3936
17.8584 | -105.3977
17.8621 | -105.4021
17.8657 | -105.4066
17.8694 | -105.4102
17.8731 | -105.4132
17.8767 | -105.4191
17.8804 | -105.426
17.8841 | -105.4322
17.8877 | -105.4396
17.8914 | -105.4477
17.8951 | -105.4557
17.8987 | -105.4629
17.9024 | -105.4705
17.9061 | -105.478
17.9097 | -105.485
17.9134 | -105.491
17.9171 | -105.4951
17.9207 | -105.4988
17.9244 | -105.5
17.9281 | -105.5
17.9317 | -105.499
17.9354 | -105.4945
17.9391 | -105.4881
17.9427 | -105.4804
17.9464 | -105.4667
17.9501 | -105.4509
17.9537 | -105.4363
17.9574 | -105.4194
17.9611 | -105.3996
17.9647 | -105.3792
17.9684 | -105.3605
17.972 | -105.3393
17.9757 | -105.3194
17.9794 | -105.2894
17.983 | -105.2666
17.9867 | -105.2429
17.9904 | -105.2182
17.994 | -105.1909
17.9977 | -105.1633
18.0014 | -105.138
18.005 | -105.1113
18.0087 | -105.0844
18.0124 | -105.0587
18.016 | -105.034
18.0197 | -105.0077
18.0234 | -104.9791
18.027 | -104.9489
18.0307 | -104.9207
18.0344 | -104.8933
18.038 | -104.8665
18.0417 | -104.8423
18.0454 | -104.8196
18.049 | -104.797
18.0527 | -104.7755
18.0564 | -104.7552
18.06 | -104.7344
18.0637 | -104.7158
18.0674 | -104.7001
18.071 | -104.6872
18.0747 | -104.6744
18.0784 | -104.662
18.082 | -104.6502
18.0857 | -104.6401
18.0893 | -104.6318
18.093 | -104.6234
18.0967 | -104.614
18.1003 | -104.6066
18.104 | -104.5989
18.1077 | -104.5909
18.1113 | -104.5843
18.115 | -104.5778
18.1187 | -104.5713
18.1223 | -104.5668
18.126 | -104.5641
18.1297 | -104.563
18.1333 | -104.5614
18.137 | -104.559
18.1407 | -104.558
18.1443 | -104.5565
18.148 | -104.5544
18.1517 | -104.5536
18.1553 | -104.553
18.159 | -104.5523
18.1627 | -104.5508
18.1663 | -104.5506
18.17 | -104.5507
18.1737 | -104.5508
18.1773 | -104.552
18.181 | -104.5537
18.1847 | -104.5569
18.1883 | -104.558
18.192 | -104.5605
18.1956 | -104.5671
18.1993 | -104.5716
18.203 | -104.5791
18.2066 | -104.5882
18.2103 | -104.597
18.214 | -104.6055
18.2176 | -104.6161
18.2213 | -104.6272
18.225 | -104.6354
18.2286 | -104.6445
18.2323 | -104.6539
18.236 | -104.6633
18.2396 | -104.6731
18.2433 | -104.683
18.247 | -104.6933
18.2506 | -104.7043
18.2543 | -104.7169
18.258 | -104.7286
18.2616 | -104.7389
18.2653 | -104.7485
18.269 | -104.758
18.2726 | -104.7664
18.2763 | -104.772
18.28 | -104.7768
18.2836 | -104.7852
18.2873 | -104.7939
18.291 | -104.8007
18.2946 | -104.8068
18.2983 | -104.8126
18.3019 | -104.8157
18.3056 | -104.8174
18.3093 | -104.8212
18.3129 | -104.8241
18.3166 | -104.8252
18.3203 | -104.828
18.3239 | -104.8307
18.3276 | -104.8319
18.3313 | -104.8338
18.3349 | -104.8362
18.3386 | -104.8397
18.3423 | -104.8437
18.3459 | -104.8497
18.3496 | -104.8556
18.3533 | -104.8634
18.3569 | -104.874
18.3606 | -104.8849
18.3643 | -104.8948
18.3679 | -104.9069
18.3716 | -104.9183
18.3753 | -104.9297
18.3789 | -104.9409
18.3826 | -104.9507
18.3863 | -104.9587
18.3899 | -104.9677
18.3936 | -104.9747
18.3973 | -104.9825
18.4009 | -104.9919
18.4046 | -105.0014
18.4082 | -105.0126
18.4119 | -105.0206
18.4156 | -105.0289
18.4192 | -105.0384
18.4229 | -105.0471
18.4266 | -105.0539
18.4302 | -105.0613
18.4339 | -105.0695
18.4376 | -105.075
18.4412 | -105.0797
18.4449 | -105.0832
18.4486 | -105.0892
18.4522 | -105.0941
18.4559 | -105.0969
18.4596 | -105.1003
18.4632 | -105.1024
18.4669 | -105.1033
18.4706 | -105.1044
18.4742 | -105.105
18.4779 | -105.1042
18.4816 | -105.1054
18.4852 | -105.1092
18.4889 | -105.1138
18.4926 | -105.1162
18.4962 | -105.1212
18.4999 | -105.1278
18.5036 | -105.132
18.5072 | -105.1387
18.5109 | -105.1442
18.5145 | -105.1497
18.5182 | -105.1555
18.5219 | -105.1592
18.5255 | -105.1637
18.5292 | -105.1704
18.5329 | -105.1761
18.5365 | -105.1823
18.5402 | -105.1894
18.5439 | -105.1978
18.5475 | -105.2056
18.5512 | -105.2129
18.5549 | -105.2191
18.5585 | -105.223
18.5622 | -105.2279
18.5659 | -105.2353
18.5695 | -105.2418
18.5732 | -105.2475
18.5769 | -105.2529
18.5805 | -105.2584
18.5842 | -105.2641
18.5879 | -105.2699
18.5915 | -105.2764
18.5952 | -105.2829
18.5989 | -105.2893
18.6025 | -105.2965
18.6062 | -105.3029
18.6099 | -105.3054
18.6135 | -105.3079
18.6172 | -105.3106
18.6208 | -105.3124
18.6245 | -105.3143
18.6282 | -105.3171
18.6318 | -105.3198
18.6355 | -105.3205
18.6392 | -105.3209
18.6428 | -105.3227
18.6465 | -105.3234
18.6502 | -105.3255
18.6538 | -105.325
18.6575 | -105.3261
18.6612 | -105.325
18.6648 | -105.3233
18.6685 | -105.3239
18.6722 | -105.3234
18.6758 | -105.3238
18.6795 | -105.3226
18.6832 | -105.3197
18.6868 | -105.3185
18.6905 | -105.3181
18.6942 | -105.3166
18.6978 | -105.3134
18.7015 | -105.31
18.7052 | -105.3072
18.7088 | -105.3052
18.7125 | -105.3031
18.7162 | -105.3002
18.7198 | -105.2989
18.7235 | -105.2984
18.7271 | -105.2974
18.7308 | -105.2972
18.7345 | -105.2965
18.7381 | -105.2971
18.7418 | -105.2991
18.7455 | -105.2996
18.7491 | -105.2987
18.7528 | -105.2992
18.7565 | -105.3031
18.7601 | -105.3041
18.7638 | -105.3054
18.7675 | -105.3085
18.7711 | -105.3089
18.7748 | -105.3076
18.7785 | -105.3097
18.7821 | -105.3127
18.7858 | -105.3147
18.7895 | -105.3185
18.7931 | -105.3212
18.7968 | -105.3222
18.8005 | -105.3243
18.8041 | -105.3271
18.8078 | -105.3316
18.8115 | -105.3346
18.8151 | -105.337
18.8188 | -105.34
18.8225 | -105.3402
18.8261 | -105.3423
18.8298 | -105.3427
18.8334 | -105.3387
18.8371 | -105.3365
18.8408 | -105.3328
18.8444 | -105.327
18.8481 | -105.3221
18.8518 | -105.3169
18.8554 | -105.3112
18.8591 | -105.3035
18.8628 | -105.2985
18.8664 | -105.2918
18.8701 | -105.2849
18.8738 | -105.2783
18.8774 | -105.2725
18.8811 | -105.267
18.8848 | -105.2616
18.8884 | -105.2572
18.8921 | -105.2499
18.8958 | -105.2426
18.8994 | -105.2361
18.9031 | -105.2316
18.9068 | -105.227
18.9104 | -105.221
18.9141 | -105.2188
18.9178 | -105.2158
18.9214 | -105.2138
18.9251 | -105.2129
18.9288 | -105.2157
18.9324 | -105.2191
18.9361 | -105.2214
18.9397 | -105.2255
18.9434 | -105.2324
18.9471 | -105.2404
18.9507 | -105.2492
18.9544 | -105.2587
18.9581 | -105.2697
18.9617 | -105.282
18.9654 | -105.2961
18.9691 | -105.3111
18.9727 | -105.3257
18.9764 | -105.34
18.9801 | -105.3552
18.9837 | -105.3708
18.9874 | -105.3871
18.9911 | -105.4048
18.9947 | -105.423
18.9984 | -105.4403
19.0021 | -105.4576
19.0057 | -105.4774
19.0094 | -105.4972
19.0131 | -105.5176
19.0167 | -105.5378
19.0204 | -105.5559
19.0241 | -105.5706
19.0277 | -105.5858
19.0314 | -105.601
19.0351 | -105.608
19.0387 | -105.6221
19.0424 | -105.6383
19.046 | -105.6538
19.0497 | -105.6632
19.0534 | -105.6719
19.057 | -105.6805
19.0607 | -105.6882
19.0644 | -105.6937
19.068 | -105.6978
19.0717 | -105.6997
19.0754 | -105.699
19.079 | -105.6962
19.0827 | -105.693
19.0864 | -105.6886
19.09 | -105.6834
19.0937 | -105.6786
19.0974 | -105.6708
19.101 | -105.6611
19.1047 | -105.6526
19.1084 | -105.6428
19.112 | -105.6322
19.1157 | -105.6226
19.1194 | -105.6104
19.123 | -105.5954
19.1267 | -105.5806
19.1304 | -105.569
19.134 | -105.5539
19.1377 | -105.5394
19.1414 | -105.5255
19.145 | -105.5119
19.1487 | -105.5028
19.1523 | -105.4923
19.156 | -105.48
19.1597 | -105.4681
19.1633 | -105.4584
19.167 | -105.4489
19.1707 | -105.439
19.1743 | -105.4297
19.178 | -105.4221
19.1817 | -105.4169
19.1853 | -105.4091
19.189 | -105.4009
19.1927 | -105.3947
19.1963 | -105.3892
19.2 | -105.3828
19.2037 | -105.3775
19.2073 | -105.3729
19.211 | -105.3689
19.2147 | -105.3661
19.2183 | -105.364
19.222 | -105.3611
19.2257 | -105.3576
19.2293 | -105.3538
19.233 | -105.3502
19.2367 | -105.3461
19.2403 | -105.3438
19.244 | -105.3412
19.2477 | -105.3371
19.2513 | -105.3346
19.255 | -105.3351
19.2586 | -105.3349
19.2623 | -105.3341
19.266 | -105.3316
19.2696 | -105.3287
19.2733 | -105.3263
19.277 | -105.3226
19.2806 | -105.3198
19.2843 | -105.3171
19.288 | -105.3147
19.2916 | -105.3127
19.2953 | -105.3129
19.299 | -105.3136
19.3026 | -105.3128
19.3063 | -105.3118
19.31 | -105.3095
19.3136 | -105.3058
19.3173 | -105.3012
19.321 | -105.2956
19.3246 | -105.2919
19.3283 | -105.287
19.332 | -105.2832
19.3356 | -105.2783
19.3393 | -105.2729
19.343 | -105.2678
19.3466 | -105.263
19.3503 | -105.2603
19.354 | -105.2564
19.3576 | -105.252
19.3613 | -105.2472
19.3649 | -105.2424
19.3686 | -105.2377
19.3723 | -105.2335
19.3759 | -105.2285
19.3796 | -105.2229
19.3833 | -105.2161
19.3869 | -105.2085
19.3906 | -105.2015
19.3943 | -105.1966
19.3979 | -105.1919
19.4016 | -105.1866
19.4053 | -105.1809
19.4089 | -105.1732
19.4126 | -105.1666
19.4163 | -105.1595
19.4199 | -105.1527
19.4236 | -105.1437
19.4273 | -105.1367
19.4309 | -105.1299
19.4346 | -105.1228
19.4383 | -105.117
19.4419 | -105.1103
19.4456 | -105.1024
19.4493 | -105.0955
19.4529 | -105.0874
19.4566 | -105.0781
19.4603 | -105.072
19.4639 | -105.0679
19.4676 | -105.0612
19.4712 | -105.0542
19.4749 | -105.0484
19.4786 | -105.0419
19.4822 | -105.036
19.4859 | -105.0308
19.4896 | -105.0255
19.4932 | -105.0219
19.4969 | -105.022
19.5006 | -105.0237
19.5042 | -105.0269
19.5079 | -105.0298
19.5116 | -105.0324
19.5152 | -105.0339
19.5189 | -105.0371
19.5226 | -105.0417
19.5262 | -105.0445
19.5299 | -105.0479
19.5336 | -105.0505
19.5372 | -105.054
19.5409 | -105.0614
19.5446 | -105.0684
19.5482 | -105.0745
19.5519 | -105.0804
19.5556 | -105.0854
19.5592 | -105.0901
19.5629 | -105.0925
19.5666 | -105.0995
19.5702 | -105.1055
19.5739 | -105.1097
19.5775 | -105.1143
19.5812 | -105.1189
19.5849 | -105.1224
19.5885 | -105.125
19.5922 | -105.1283
19.5959 | -105.1323
19.5995 | -105.1375
19.6032 | -105.1411
19.6069 | -105.1451
19.6105 | -105.1462
19.6142 | -105.1478
19.6179 | -105.1498
19.6215 | -105.1516
19.6252 | -105.1495
19.6289 | -105.1479
19.6325 | -105.1472
19.6362 | -105.1446
19.6399 | -105.1438
19.6435 | -105.1436
19.6472 | -105.1448
19.6509 | -105.1456
19.6545 | -105.1471
19.6582 | -105.1492
19.6619 | -105.153
19.6655 | -105.1576
19.6692 | -105.1603
19.6729 | -105.1615
19.6765 | -105.1647
19.6802 | -105.1688
19.6838 | -105.1712
19.6875 | -105.1769
19.6912 | -105.1847
19.6948 | -105.1941
19.6985 | -105.2029
19.7022 | -105.2118
19.7058 | -105.2188
19.7095 | -105.2265
19.7132 | -105.2343
19.7168 | -105.2414
19.7205 | -105.2521
19.7242 | -105.2615
19.7278 | -105.2675
19.7315 | -105.2737
19.7352 | -105.2834
19.7388 | -105.2915
19.7425 | -105.298
19.7462 | -105.3056
19.7498 | -105.3117
19.7535 | -105.3174
19.7572 | -105.3227
19.7608 | -105.3292
19.7645 | -105.3366
19.7682 | -105.3432
19.7718 | -105.3479
19.7755 | -105.3496
19.7792 | -105.3513
19.7828 | -105.3513
19.7865 | -105.3523
19.7901 | -105.3543
19.7938 | -105.3562
19.7975 | -105.3561
19.8011 | -105.3572
19.8048 | -105.3592
19.8085 | -105.3576
19.8121 | -105.3574
19.8158 | -105.3568
19.8195 | -105.3544
19.8231 | -105.3522
19.8268 | -105.3504
19.8305 | -105.3463
19.8341 | -105.3411
19.8378 | -105.3388
19.8415 | -105.3347
19.8451 | -105.3288
19.8488 | -105.3217
19.8525 | -105.3147
19.8561 | -105.3076
19.8598 | -105.3008
19.8635 | -105.2947
19.8671 | -105.2899
19.8708 | -105.2873
19.8745 | -105.2831
19.8781 | -105.2783
19.8818 | -105.2721
19.8855 | -105.2672
19.8891 | -105.2613
19.8928 | -105.2569
19.8964 | -105.2532
19.9001 | -105.2494
19.9038 | -105.2453
19.9074 | -105.2442
19.9111 | -105.2391
19.9148 | -105.2349
19.9184 | -105.2315
19.9221 | -105.2265
19.9258 | -105.222
19.9294 | -105.2168
19.9331 | -105.2124
19.9368 | -105.2091
19.9404 | -105.2044
19.9441 | -105.1978
19.9478 | -105.1895
19.9514 | -105.1839
19.9551 | -105.1784
19.9588 | -105.1713
19.9624 | -105.1649
19.9661 | -105.1617
19.9698 | -105.1567
19.9734 | -105.1503
19.9771 | -105.1435
19.9808 | -105.1371
19.9844 | -105.1321
19.9881 | -105.1288
19.9918 | -105.1266
19.9954 | -105.1242
19.9991 | -105.1211
20.0027 | -105.1175
20.0064 | -105.1176
20.0101 | -105.1187
20.0137 | -105.119
20.0174 | -105.1163
20.0211 | -105.115
20.0247 | -105.1157
20.0284 | -105.1172
20.0321 | -105.1223
20.0357 | -105.1275
20.0394 | -105.1333
20.0431 | -105.1374
20.0467 | -105.1431
20.0504 | -105.1496
20.0541 | -105.1568
20.0577 | -105.1658
20.0614 | -105.1744
20.0651 | -105.1844
20.0687 | -105.194
20.0724 | -105.2056
20.0761 | -105.2157
20.0797 | -105.2246
20.0834 | -105.2338
20.0871 | -105.2416
20.0907 | -105.2506
20.0944 | -105.2582
20.0981 | -105.2668
20.1017 | -105.2744
20.1054 | -105.2815
20.109 | -105.2896
20.1127 | -105.2977
20.1164 | -105.307
20.12 | -105.3149
20.1237 | -105.32
20.1274 | -105.3285
20.131 | -105.3341
20.1347 | -105.3396
20.1384 | -105.3455
20.142 | -105.3521
20.1457 | -105.3574
20.1494 | -105.3616
20.153 | -105.3662
20.1567 | -105.368
20.1604 | -105.3708
20.164 | -105.3747
20.1677 | -105.3803
20.1714 | -105.3863
20.175 | -105.3936
20.1787 | -105.3996
20.1824 | -105.4057
20.186 | -105.4122
20.1897 | -105.4173
20.1934 | -105.421
20.197 | -105.4258
20.2007 | -105.4309
20.2044 | -105.4353
20.208 | -105.4402
20.2117 | -105.4415
20.2153 | -105.4438
20.219 | -105.448
20.2227 | -105.4495
20.2263 | -105.4502
20.23 | -105.4522
20.2337 | -105.4528
20.2373 | -105.4508
20.241 | -105.4471
20.2447 | -105.4438
20.2483 | -105.4381
20.252 | -105.4309
20.2557 | -105.4236
20.2593 | -105.4146
20.263 | -105.4036
20.2667 | -105.3934
20.2703 | -105.3815
20.274 | -105.3671
20.2777 | -105.3524
20.2813 | -105.336
20.285 | -105.3198
20.2887 | -105.2998
20.2923 | -105.2778
20.296 | -105.2564
20.2997 | -105.2332
20.3033 | -105.2098
20.307 | -105.1844
20.3107 | -105.1565
20.3143 | -105.1273
20.318 | -105.0982
20.3216 | -105.071
20.3253 | -105.0433
20.329 | -105.0157
20.3326 | -104.9865
20.3363 | -104.957
20.34 | -104.9289
20.3436 | -104.9027
20.3473 | -104.8766
20.351 | -104.8505
20.3546 | -104.825
20.3583 | -104.7984
20.362 | -104.7754
20.3656 | -104.7563
20.3693 | -104.7376
20.373 | -104.7189
20.3766 | -104.7028
20.3803 | -104.6893
20.384 | -104.6763
20.3876 | -104.6652
20.3913 | -104.6564
20.395 | -104.6502
20.3986 | -104.6462
20.4023 | -104.6445
20.406 | -104.6453
20.4096 | -104.6496
20.4133 | -104.6544
20.417 | -104.6582
20.4206 | -104.6669
20.4243 | -104.6768
20.4279 | -104.6873
20.4316 | -104.6981
20.4353 | -104.7116
20.4389 | -104.7265
20.4426 | -104.7416
20.4463 | -104.7571
20.4499 | -104.7744
20.4536 | -104.7919
20.4573 | -104.8095
20.4609 | -104.825
20.4646 | -104.8397
20.4683 | -104.8576
20.4719 | -104.8746
20.4756 | -104.8901
20.4793 | -104.9053
20.4829 | -104.9201
20.4866 | -104.9323
20.4903 | -104.9416
20.4939 | -104.9502
20.4976 | -104.9554
20.5013 | -104.9589
20.5049 | -104.9606
20.5086 | -104.9601
20.5123 | -104.9575
20.5159 | -104.9544
20.5196 | -104.9482
20.5233 | -104.9404
20.5269 | -104.9299
20.5306 | -104.9186
20.5343 | -104.9044
20.5379 | -104.889
20.5416 | -104.8704
20.5452 | -104.8495
20.5489 | -104.828
20.5526 | -104.8085
20.5562 | -104.7871
20.5599 | -104.7631
20.5636 | -104.7405
20.5672 | -104.7167
20.5709 | -104.6912
20.5746 | -104.6652
20.5782 | -104.6413
20.5819 | -104.6174
20.5856 | -104.5933
20.5892 | -104.5716
20.5929 | -104.5504
20.5966 | -104.5298
20.6002 | -104.5096
20.6039 | -104.4894
20.6076 | -104.4682
20.6112 | -104.4489
20.6149 | -104.4303
20.6186 | -104.4118
20.6222 | -104.4046
20.6259 | -104.3887
20.6296 | -104.3761
20.6332 | -104.3664
20.6369 | -104.3553
20.6406 | -104.345
20.6442 | -104.3346
20.6479 | -104.3232
20.6515 | -104.3145
20.6552 | -104.3084
20.6589 | -104.3019
20.6625 | -104.2968
20.6662 | -104.2947
20.6699 | -104.2926
20.6735 | -104.2897
20.6772 | -104.2875
20.6809 | -104.288
20.6845 | -104.2898
20.6882 | -104.292
20.6919 | -104.2932
20.6955 | -104.2943
20.6992 | -104.2962
20.7029 | -104.2974
20.7065 | -104.3015
20.7102 | -104.3066
20.7139 | -104.3106
20.7175 | -104.3144
20.7212 | -104.3192
20.7249 | -104.3233
20.7285 | -104.3258
20.7322 | -104.3286
20.7359 | -104.3302
20.7395 | -104.332
20.7432 | -104.3334
20.7468 | -104.3339
20.7505 | -104.3316
20.7542 | -104.3304
20.7578 | -104.3305
20.7615 | -104.3296
20.7652 | -104.3281
20.7688 | -104.3276
20.7725 | -104.3274
20.7762 | -104.3254
20.7798 | -104.3233
20.7835 | -104.3224
20.7872 | -104.3209
20.7908 | -104.3191
20.7945 | -104.319
20.7982 | -104.3183
20.8018 | -104.3166
20.8055 | -104.3153
20.8092 | -104.3132
20.8128 | -104.3122
20.8165 | -104.3121
20.8202 | -104.31
20.8238 | -104.3073
20.8275 | -104.3056
20.8312 | -104.3048
20.8348 | -104.3045
20.8385 | -104.3055
20.8422 | -104.3075
20.8458 | -104.3083
20.8495 | -104.3087
20.8531 | -104.3101
20.8568 | -104.3119
20.8605 | -104.3145
20.8641 | -104.3195
20.8678 | -104.3241
20.8715 | -104.3271
20.8751 | -104.3304
20.8788 | -104.3341
20.8825 | -104.3392
20.8861 | -104.3444
20.8898 | -104.3511
20.8935 | -104.3588
20.8971 | -104.3664
20.9008 | -104.376
20.9045 | -104.3846
20.9081 | -104.3948
20.9118 | -104.4074
20.9155 | -104.4193
20.9191 | -104.4309
20.9228 | -104.4436
20.9265 | -104.456
20.9301 | -104.4655
20.9338 | -104.4766
20.9375 | -104.4903
20.9411 | -104.5067
20.9448 | -104.525
20.9485 | -104.5427
20.9521 | -104.5577
20.9558 | -104.5728
20.9595 | -104.5862
20.9631 | -104.5998
20.9668 | -104.6143
20.9704 | -104.629
20.9741 | -104.6441
20.9778 | -104.6572
20.9814 | -104.6695
20.9851 | -104.6816
20.9888 | -104.695
20.9924 | -104.7073
20.9961 | -104.719
20.9998 | -104.7313
21.0034 | -104.7418
21.0071 | -104.7519
21.0108 | -104.7629
21.0144 | -104.7731
21.0181 | -104.7822
21.0218 | -104.7918
21.0254 | -104.7998
21.0291 | -104.8077
21.0328 | -104.8161
21.0364 | -104.8237
21.0401 | -104.8287
21.0438 | -104.8307
21.0474 | -104.8335
21.0511 | -104.8369
21.0548 | -104.8397
21.0584 | -104.8395
21.0621 | -104.8382
21.0658 | -104.8365
21.0694 | -104.8338
21.0731 | -104.8323
21.0767 | -104.8307
21.0804 | -104.825
21.0841 | -104.8154
21.0877 | -104.8072
21.0914 | -104.7977
21.0951 | -104.7877
21.0987 | -104.778
21.1024 | -104.7673
21.1061 | -104.757
21.1097 | -104.7462
21.1134 | -104.7332
21.1171 | -104.721
21.1207 | -104.7089
21.1244 | -104.6947
21.1281 | -104.6775
21.1317 | -104.6589
21.1354 | -104.6395
21.1391 | -104.6193
21.1427 | -104.5988
21.1464 | -104.5791
21.1501 | -104.5563
21.1537 | -104.5361
21.1574 | -104.5162
21.1611 | -104.496
21.1647 | -104.4766
21.1684 | -104.4568
21.172 | -104.4372
21.1757 | -104.4163
21.1794 | -104.3951
21.183 | -104.3746
21.1867 | -104.3565
21.1904 | -104.3396
21.194 | -104.323
21.1977 | -104.3081
21.2014 | -104.2916
21.205 | -104.276
21.2087 | -104.2632
21.2124 | -104.2513
21.216 | -104.2411
21.2197 | -104.2327
21.2234 | -104.2244
21.227 | -104.2174
21.2307 | -104.2144
21.2344 | -104.2121
21.238 | -104.2087
21.2417 | -104.207
21.2454 | -104.2061
21.249 | -104.2034
21.2527 | -104.2019
21.2564 | -104.2033
21.26 | -104.2068
21.2637 | -104.2103
21.2674 | -104.2142
21.271 | -104.218
21.2747 | -104.2214
21.2784 | -104.2247
21.282 | -104.23
21.2857 | -104.235
21.2893 | -104.2393
21.293 | -104.243
21.2967 | -104.2466
21.3003 | -104.2509
21.304 | -104.2559
21.3077 | -104.26
21.3113 | -104.2633
21.315 | -104.266
21.3187 | -104.2695
21.3223 | -104.2717
21.326 | -104.2731
21.3297 | -104.2763
21.3333 | -104.2807
21.337 | -104.2855
21.3407 | -104.2902
21.3443 | -104.2962
21.348 | -104.3023
21.3517 | -104.3078
21.3553 | -104.31
21.359 | -104.3118
21.3627 | -104.3159
21.3663 | -104.3204
21.37 | -104.3267
21.3737 | -104.3329
21.3773 | -104.3374
21.381 | -104.3423
21.3847 | -104.3491
21.3883 | -104.3544
21.392 | -104.3579
21.3956 | -104.36
21.3993 | -104.3609
21.403 | -104.3635
21.4066 | -104.3657
21.4103 | -104.3671
21.414 | -104.3686
21.4176 | -104.3699
21.4213 | -104.3699
21.425 | -104.371
21.4286 | -104.3724
21.4323 | -104.3733
21.436 | -104.3745
21.4396 | -104.3737
21.4433 | -104.3727
21.447 | -104.3721
21.4506 | -104.3732
21.4543 | -104.3754
21.458 | -104.3782
21.4616 | -104.3786
21.4653 | -104.3805
21.469 | -104.3828
21.4726 | -104.3862
21.4763 | -104.392
21.48 | -104.3984
21.4836 | -104.4047
21.4873 | -104.4101
21.491 | -104.4121
21.4946 | -104.4121
21.4983 | -104.4122
21.5019 | -104.4169
21.5056 | -104.4243
21.5093 | -104.4328
21.5129 | -104.439
21.5166 | -104.4436
21.5203 | -104.4466
21.5239 | -104.4483
21.5276 | -104.451
21.5313 | -104.4526
21.5349 | -104.4505
21.5386 | -104.4464
21.5423 | -104.442
21.5459 | -104.4366
21.5496 | -104.4312
21.5533 | -104.426
21.5569 | -104.4178
21.5606 | -104.4044
21.5643 | -104.3906
21.5679 | -104.376
21.5716 | -104.3587
21.5753 | -104.3393
21.5789 | -104.3199
21.5826 | -104.2996
21.5863 | -104.2764
21.5899 | -104.252
21.5936 | -104.2264
21.5973 | -104.2012
21.6009 | -104.1732
21.6046 | -104.1427
21.6082 | -104.1101
21.6119 | -104.0766
21.6156 | -104.0425
21.6192 | -104.0082
21.6229 | -103.974
21.6266 | -103.9401
21.6302 | -103.9057
21.6339 | -103.8702
21.6376 | -103.8349
21.6412 | -103.8001
21.6449 | -103.7648
21.6486 | -103.7314
21.6522 | -103.6988
21.6559 | -103.6669
21.6596 | -103.6389
21.6632 | -103.6135
21.6669 | -103.5883
21.6706 | -103.5652
21.6742 | -103.5425
21.6779 | -103.5227
21.6816 | -103.5257
21.6852 | -103.5124
21.6889 | -103.5027
21.6926 | -103.4967
21.6962 | -103.4915
21.6999 | -103.489
21.7036 | -103.4892
21.7072 | -103.4933
21.7109 | -103.4987
21.7145 | -103.5065
21.7182 | -103.5175
21.7219 | -103.5314
21.7255 | -103.5483
21.7292 | -103.5675
21.7329 | -103.5886
21.7365 | -103.6111
21.7402 | -103.6361
21.7439 | -103.6622
21.7475 | -103.6894
21.7512 | -103.7183
21.7549 | -103.7492
21.7585 | -103.7808
21.7622 | -103.8132
21.7659 | -103.8464
21.7695 | -103.881
21.7732 | -103.9166
21.7769 | -103.9519
21.7805 | -103.987
21.7842 | -104.0201
21.7879 | -104.0533
21.7915 | -104.0877
21.7952 | -104.1205
21.7989 | -104.1554
21.8025 | -104.19
21.8062 | -104.2218
21.8099 | -104.248
21.8135 | -104.2704
21.8172 | -104.2902
21.8208 | -104.3109
21.8245 | -104.3327
21.8282 | -104.3515
21.8318 | -104.3684
21.8355 | -104.3828
21.8392 | -104.3928
21.8428 | -104.4012
21.8465 | -104.4104
21.8502 | -104.4193
21.8538 | -104.4244
21.8575 | -104.4252
21.8612 | -104.4228
21.8648 | -104.4202
21.8685 | -104.4168
21.8722 | -104.4092
21.8758 | -104.4003
21.8795 | -104.3907
21.8832 | -104.3796
21.8868 | -104.3645
21.8905 | -104.3474
21.8942 | -104.3302
21.8978 | -104.3123
21.9015 | -104.2923
21.9052 | -104.2728
21.9088 | -104.2526
21.9125 | -104.2308
21.9162 | -104.2091
21.9198 | -104.1878
21.9235 | -104.1666
21.9271 | -104.1456
21.9308 | -104.1233
21.9345 | -104.0998
21.9381 | -104.0779
21.9418 | -104.0535
21.9455 | -104.0282
21.9491 | -104.0066
21.9528 | -103.9885
21.9565 | -103.9691
21.9601 | -103.9501
21.9638 | -103.9315
21.9675 | -103.9124
21.9711 | -103.8944
21.9748 | -103.8779
21.9785 | -103.8602
21.9821 | -103.8427
21.9858 | -103.827
21.9895 | -103.8149
21.9931 | -103.8035
21.9968 | -103.7924
22.0005 | -103.7825
22.0041 | -103.7718
22.0078 | -103.7627
22.0115 | -103.7567
22.0151 | -103.7517
22.0188 | -103.7473
22.0225 | -103.744
22.0261 | -103.7428
22.0298 | -103.7432
22.0334 | -103.7428
22.0371 | -103.744
22.0408 | -103.748
22.0444 | -103.7525
22.0481 | -103.7603
22.0518 | -103.7688
22.0554 | -103.7777
22.0591 | -103.7887
22.0628 | -103.8006
22.0664 | -103.8128
22.0701 | -103.8256
22.0738 | -103.8398
22.0774 | -103.8542
22.0811 | -103.8704
22.0848 | -103.8855
22.0884 | -103.9017
22.0921 | -103.9202
22.0958 | -103.9398
22.0994 | -103.9589
22.1031 | -103.9762
22.1068 | -103.9972
22.1104 | -104.0181
22.1141 | -104.0394
22.1178 | -104.0628
22.1214 | -104.0863
22.1251 | -104.1112
22.1288 | -104.1362
22.1324 | -104.1578
22.1361 | -104.1787
22.1397 | -104.2001
22.1434 | -104.2202
22.1471 | -104.241
22.1507 | -104.263
22.1544 | -104.284
22.1581 | -104.3026
22.1617 | -104.3185
22.1654 | -104.335
22.1691 | -104.3519
22.1727 | -104.3689
22.1764 | -104.3851
22.1801 | -104.4017
22.1837 | -104.4164
22.1874 | -104.4265
22.1911 | -104.4355
22.1947 | -104.4456
22.1984 | -104.4561
22.2021 | -104.4651
22.2057 | -104.4721
22.2094 | -104.472
22.2131 | -104.4771
22.2167 | -104.4822
22.2204 | -104.4837
22.2241 | -104.488
22.2277 | -104.4913
22.2314 | -104.4913
22.2351 | -104.4918
22.2387 | -104.4914
22.2424 | -104.4883
22.246 | -104.4847
22.2497 | -104.4827
22.2534 | -104.4809
22.257 | -104.4798
22.2607 | -104.476
22.2644 | -104.4679
22.268 | -104.46
22.2717 | -104.4536
22.2754 | -104.4454
22.279 | -104.4348
22.2827 | -104.4241
22.2864 | -104.4136
22.29 | -104.4032
22.2937 | -104.3937
22.2974 | -104.3848
22.301 | -104.3734
22.3047 | -104.3583
22.3084 | -104.3439
22.312 | -104.33
22.3157 | -104.3151
22.3194 | -104.2993
22.323 | -104.284
22.3267 | -104.267
22.3304 | -104.2456
22.334 | -104.2251
22.3377 | -104.2067
22.3414 | -104.1865
22.345 | -104.165
22.3487 | -104.1463
22.3523 | -104.1268
22.356 | -104.1051
22.3597 | -104.0821
22.3633 | -104.0582
22.367 | -104.0333
22.3707 | -104.0087
22.3743 | -103.9833
22.378 | -103.9568
22.3817 | -103.93
22.3853 | -103.9017
22.389 | -103.8717
22.3927 | -103.8444
22.3963 | -103.8165
22.4 | -103.7879
22.4037 | -103.7591
22.4073 | -103.7284
22.411 | -103.698
22.4147 | -103.6694
22.4183 | -103.6411
22.422 | -103.609
22.4257 | -103.5759
22.4293 | -103.5426
22.433 | -103.5091
22.4367 | -103.4776
22.4403 | -103.4461
22.444 | -103.4157
22.4477 | -103.3882
22.4513 | -103.3608
22.455 | -103.3363
22.4586 | -103.3132
22.4623 | -103.2897
22.466 | -103.2653
22.4696 | -103.2427
22.4733 | -103.2201
22.477 | -103.1984
22.4806 | -103.1782
22.4843 | -103.1598
22.488 | -103.143
22.4916 | -103.1265
22.4953 | -103.1107
22.499 | -103.0972
22.5026 | -103.0846
22.5063 | -103.073
22.51 | -103.0629
22.5136 | -103.0544
22.5173 | -103.048
22.521 | -103.0417
22.5246 | -103.0359
22.5283 | -103.032
22.532 | -103.0295
22.5356 | -103.0281
22.5393 | -103.0281
22.543 | -103.0275
22.5466 | -103.0287
22.5503 | -103.0332
22.554 | -103.0345
22.5576 | -103.0355
22.5613 | -103.0393
22.5649 | -103.0457
22.5686 | -103.0528
22.5723 | -103.0601
22.5759 | -103.0674
22.5796 | -103.0784
22.5833 | -103.0898
22.5869 | -103.0994
22.5906 | -103.1093
22.5943 | -103.1208
22.5979 | -103.1343
22.6016 | -103.1487
22.6053 | -103.1616
22.6089 | -103.1727
22.6126 | -103.1822
22.6163 | -103.1936
22.6199 | -103.2056
22.6236 | -103.2163
22.6273 | -103.2264
22.6309 | -103.2384
22.6346 | -103.2487
22.6383 | -103.2605
22.6419 | -103.2719
22.6456 | -103.2824
22.6493 | -103.2944
22.6529 | -103.3034
22.6566 | -103.3118
22.6603 | -103.3214
22.6639 | -103.3311
22.6676 | -103.3402
22.6712 | -103.3496
22.6749 | -103.3584
22.6786 | -103.369
22.6822 | -103.3811
22.6859 | -103.3958
22.6896 | -103.4098
22.6932 | -103.4232
22.6969 | -103.4379
22.7006 | -103.4525
22.7042 | -103.4685
22.7079 | -103.4855
22.7116 | -103.5004
22.7152 | -103.5157
22.7189 | -103.5328
22.7226 | -103.5523
22.7262 | -103.5722
22.7299 | -103.5928
22.7336 | -103.6132
22.7372 | -103.6377
22.7409 | -103.6591
22.7446 | -103.6821
22.7482 | -103.7058
22.7519 | -103.7299
22.7556 | -103.7538
22.7592 | -103.7759
22.7629 | -103.7968
22.7666 | -103.8158
22.7702 | -103.8362
22.7739 | -103.8567
22.7775 | -103.8742
22.7812 | -103.8904
22.7849 | -103.9078
22.7885 | -103.9272
22.7922 | -103.945
22.7959 | -103.9615
22.7995 | -103.9782
22.8032 | -103.9937
22.8069 | -104.0076
22.8105 | -104.0195
22.8142 | -104.0323
22.8179 | -104.0447
22.8215 | -104.0559
22.8252 | -104.0688
22.8289 | -104.0776
22.8325 | -104.0833
22.8362 | -104.0863
22.8399 | -104.0894
22.8435 | -104.0935
22.8472 | -104.0981
22.8509 | -104.1
22.8545 | -104.1008
22.8582 | -104.1015
22.8619 | -104.101
22.8655 | -104.0995
22.8692 | -104.0989
22.8729 | -104.097
22.8765 | -104.0968
22.8802 | -104.0947
22.8838 | -104.0886
22.8875 | -104.0824
22.8912 | -104.0746
22.8948 | -104.0667
22.8985 | -104.0601
22.9022 | -104.0508
22.9058 | -104.0424
22.9095 | -104.035
22.9132 | -104.0269
22.9168 | -104.0191
22.9205 | -104.0111
22.9242 | -104.0025
22.9278 | -103.9942
22.9315 | -103.9871
22.9352 | -103.982
22.9388 | -103.9753
22.9425 | -103.9686
22.9462 | -103.964
22.9498 | -103.961
22.9535 | -103.9571
22.9572 | -103.9545
22.9608 | -103.9514
22.9645 | -103.9484
22.9682 | -103.9467
22.9718 | -103.9466
22.9755 | -103.9472
22.9792 | -103.9453
22.9828 | -103.9451
22.9865 | -103.9459
22.9901 | -103.9486
22.9938 | -103.951
22.9975 | -103.9547
23.0011 | -103.958
23.0048 | -103.9603
23.0085 | -103.9607
23.0121 | -103.9634
23.0158 | -103.9681
23.0195 | -103.9717
23.0231 | -103.9736
23.0268 | -103.9784
23.0305 | -103.9851
23.0341 | -103.9908
23.0378 | -103.9968
23.0415 | -104.0018
23.0451 | -104.0029
23.0488 | -104.0049
23.0525 | -104.0081
23.0561 | -104.0097
23.0598 | -104.0092
23.0635 | -104.0095
23.0671 | -104.01
23.0708 | -104.0092
23.0745 | -104.0078
23.0781 | -104.0063
23.0818 | -104.0042
23.0855 | -104.0027
23.0891 | -104.0018
23.0928 | -104.0005
23.0964 | -103.9926
23.1001 | -103.9833
23.1038 | -103.9774
23.1074 | -103.9717
23.1111 | -103.9632
23.1148 | -103.9527
23.1184 | -103.9426
23.1221 | -103.9324
23.1258 | -103.9221
23.1294 | -103.9119
23.1331 | -103.9007
23.1368 | -103.8904
23.1404 | -103.8783
23.1441 | -103.8637
23.1478 | -103.8501
23.1514 | -103.8366
23.1551 | -103.8254
23.1588 | -103.8145
23.1624 | -103.8051
23.1661 | -103.7909
23.1698 | -103.7758
23.1734 | -103.7604
23.1771 | -103.7455
23.1808 | -103.7315
23.1844 | -103.7176
23.1881 | -103.7053
23.1918 | -103.6913
23.1954 | -103.6766
23.1991 | -103.6611
23.2027 | -103.6459
23.2064 | -103.6334
23.2101 | -103.6233
23.2137 | -103.61
23.2174 | -103.5979
23.2211 | -103.5863
23.2247 | -103.5754
23.2284 | -103.5653
23.2321 | -103.5545
23.2357 | -103.5449
23.2394 | -103.5362
23.2431 | -103.5266
23.2467 | -103.5158
23.2504 | -103.5034
23.2541 | -103.492
23.2577 | -103.4817
23.2614 | -103.4719
23.2651 | -103.4624
23.2687 | -103.4533
23.2724 | -103.4449
23.2761 | -103.4361
23.2797 | -103.4267
23.2834 | -103.4159
23.2871 | -103.4061
23.2907 | -103.3975
23.2944 | -103.3895
23.2981 | -103.382
23.3017 | -103.3736
23.3054 | -103.3672
23.309 | -103.3636
23.3127 | -103.358
23.3164 | -103.3528
23.32 | -103.3486
23.3237 | -103.346
23.3274 | -103.3436
23.331 | -103.3429
23.3347 | -103.3439
23.3384 | -103.3474
23.342 | -103.3522
23.3457 | -103.3562
23.3494 | -103.362
23.353 | -103.371
23.3567 | -103.3777
23.3604 | -103.3845
23.364 | -103.3933
23.3677 | -103.4046
23.3714 | -103.4176
23.375 | -103.4338
23.3787 | -103.4508
23.3824 | -103.467
23.386 | -103.4849
23.3897 | -103.5052
23.3934 | -103.5253
23.397 | -103.5454
23.4007 | -103.5677
23.4044 | -103.5905
23.408 | -103.6137
23.4117 | -103.6372
23.4153 | -103.6609
23.419 | -103.6846
23.4227 | -103.7093
23.4263 | -103.7378
23.43 | -103.7666
23.4337 | -103.7921
23.4373 | -103.8171
23.441 | -103.8439
23.4447 | -103.8708
23.4483 | -103.8975
23.452 | -103.9247
23.4557 | -103.9481
23.4593 | -103.9698
23.463 | -103.9916
23.4667 | -104.0129
23.4703 | -104.0344
23.474 | -104.0572
23.4777 | -104.0796
23.4813 | -104.1001
23.485 | -104.1201
23.4887 | -104.1384
23.4923 | -104.1539
23.496 | -104.1722
23.4997 | -104.1894
23.5033 | -104.204
23.507 | -104.2173
23.5107 | -104.2284
23.5143 | -104.2375
23.518 | -104.2465
23.5216 | -104.2554
23.5253 | -104.2621
23.529 | -104.2681
23.5326 | -104.2737
23.5363 | -104.2765
23.54 | -104.2752
23.5436 | -104.2754
23.5473 | -104.2754
23.551 | -104.2738
23.5546 | -104.2692
23.5583 | -104.2656
23.562 | -104.2625
23.5656 | -104.2583
23.5693 | -104.2522
23.573 | -104.2466
23.5766 | -104.2442
23.5803 | -104.2429
23.584 | -104.2399
23.5876 | -104.2351
23.5913 | -104.2277
23.595 | -104.2202
23.5986 | -104.2144
23.6023 | -104.2077
23.606 | -104.2005
23.6096 | -104.1954
23.6133 | -104.1897
23.617 | -104.1842
23.6206 | -104.1806
23.6243 | -104.1761
23.6279 | -104.173
23.6316 | -104.1713
23.6353 | -104.1709
23.6389 | -104.1691
23.6426 | -104.1674
23.6463 | -104.1667
23.6499 | -104.1663
23.6536 | -104.1659
23.6573 | -104.1656
23.6609 | -104.167
23.6646 | -104.1704
23.6683 | -104.1744
23.6719 | -104.1784
23.6756 | -104.1827
23.6793 | -104.1879
23.6829 | -104.1922
23.6866 | -104.1983
23.6903 | -104.2059
23.6939 | -104.2146
23.6976 | -104.2243
23.7013 | -104.2329
23.7049 | -104.2412
23.7086 | -104.249
23.7123 | -104.2568
23.7159 | -104.2659
23.7196 | -104.2765
23.7233 | -104.2873
23.7269 | -104.2976
23.7306 | -104.3089
23.7342 | -104.3211
23.7379 | -104.3338
23.7416 | -104.3469
23.7452 | -104.36
23.7489 | -104.3723
23.7526 | -104.3828
23.7562 | -104.394
23.7599 | -104.4063
23.7636 | -104.4188
23.7672 | -104.4313
23.7709 | -104.4439
23.7746 | -104.4541
23.7782 | -104.4642
23.7819 | -104.4747
23.7856 | -104.4843
23.7892 | -104.4964
23.7929 | -104.5097
23.7966 | -104.5193
23.8002 | -104.5316
23.8039 | -104.543
23.8076 | -104.5527
23.8112 | -104.5611
23.8149 | -104.5658
23.8186 | -104.5727
23.8222 | -104.5797
23.8259 | -104.5879
23.8296 | -104.5956
23.8332 | -104.6025
23.8369 | -104.609
23.8406 | -104.6183
23.8442 | -104.6264
23.8479 | -104.6349
23.8515 | -104.6432
23.8552 | -104.6494
23.8589 | -104.6561
23.8625 | -104.6628
23.8662 | -104.668
23.8699 | -104.6718
23.8735 | -104.6761
23.8772 | -104.6822
23.8809 | -104.6854
23.8845 | -104.6864
23.8882 | -104.6873
23.8919 | -104.6912
23.8955 | -104.6925
23.8992 | -104.6922
23.9029 | -104.6922
23.9065 | -104.6941
23.9102 | -104.6953
23.9139 | -104.693
23.9175 | -104.6914
23.9212 | -104.6904
23.9249 | -104.6901
23.9285 | -104.6881
23.9322 | -104.6866
23.9359 | -104.6838
23.9395 | -104.6829
23.9432 | -104.6813
23.9469 | -104.6778
23.9505 | -104.6744
23.9542 | -104.6731
23.9578 | -104.671
23.9615 | -104.6685
23.9652 | -104.6643
23.9688 | -104.6585
23.9725 | -104.6518
23.9762 | -104.6453
23.9798 | -104.6403
23.9835 | -104.6371
23.9872 | -104.6342
23.9908 | -104.6306
23.9945 | -104.6255
23.9982 | -104.6218
24.0018 | -104.6178
24.0055 | -104.6131
24.0092 | -104.6077
24.0128 | -104.6008
24.0165 | -104.597
24.0202 | -104.5956
24.0238 | -104.5919
24.0275 | -104.5874
24.0312 | -104.5811
24.0348 | -104.5751
24.0385 | -104.5703
24.0422 | -104.5681
24.0458 | -104.5669
24.0495 | -104.5654
24.0531 | -104.5622
24.0568 | -104.5576
24.0605 | -104.5525
24.0641 | -104.5476
24.0678 | -104.542
24.0715 | -104.5356
24.0751 | -104.5301
24.0788 | -104.5262
24.0825 | -104.5231
24.0861 | -104.5197
24.0898 | -104.5186
24.0935 | -104.5161
24.0971 | -104.5125
24.1008 | -104.51
24.1045 | -104.5083
24.1081 | -104.507
24.1118 | -104.5066
24.1155 | -104.5063
24.1191 | -104.5058
24.1228 | -104.5056
24.1265 | -104.505
24.1301 | -104.5026
24.1338 | -104.5004
24.1375 | -104.5003
24.1411 | -104.4984
24.1448 | -104.4975
24.1485 | -104.4965
24.1521 | -104.4967
24.1558 | -104.4992
24.1594 | -104.5002
24.1631 | -104.5003
24.1668 | -104.5014
24.1704 | -104.5025
24.1741 | -104.5024
24.1778 | -104.5034
24.1814 | -104.5016
24.1851 | -104.5008
24.1888 | -104.5002
24.1924 | -104.5009
24.1961 | -104.5004
24.1998 | -104.499
24.2034 | -104.4989
24.2071 | -104.4999
24.2108 | -104.5019
24.2144 | -104.5039
24.2181 | -104.5032
24.2218 | -104.5039
24.2254 | -104.5055
24.2291 | -104.507
24.2328 | -104.5063
24.2364 | -104.5051
24.2401 | -104.5048
24.2438 | -104.5046
24.2474 | -104.5018
24.2511 | -104.4993
24.2548 | -104.4964
24.2584 | -104.4914
24.2621 | -104.4842
24.2658 | -104.4754
24.2694 | -104.4677
24.2731 | -104.4576
24.2767 | -104.447
24.2804 | -104.432
24.2841 | -104.4157
24.2877 | -104.4013
24.2914 | -104.39
24.2951 | -104.3784
24.2987 | -104.3649
24.3024 | -104.3513
24.3061 | -104.3375
24.3097 | -104.3233
24.3134 | -104.3104
24.3171 | -104.2971
24.3207 | -104.2807
24.3244 | -104.26
24.3281 | -104.241
24.3317 | -104.2207
24.3354 | -104.2018
24.3391 | -104.1851
24.3427 | -104.1671
24.3464 | -104.1479
24.3501 | -104.1287
24.3537 | -104.109
24.3574 | -104.0898
24.3611 | -104.0707
24.3647 | -104.0505
24.3684 | -104.031
24.3721 | -104.0126
24.3757 | -103.9963
24.3794 | -103.9805
24.383 | -103.964
24.3867 | -103.9461
24.3904 | -103.9284
24.394 | -103.9121
24.3977 | -103.8986
24.4014 | -103.8875
24.405 | -103.8779
24.4087 | -103.8679
24.4124 | -103.8577
24.416 | -103.85
24.4197 | -103.8429
24.4234 | -103.8362
24.427 | -103.831
24.4307 | -103.8269
24.4344 | -103.8236
24.438 | -103.8236
24.4417 | -103.8231
24.4454 | -103.8239
24.449 | -103.8258
24.4527 | -103.8294
24.4564 | -103.8342
24.46 | -103.8379
24.4637 | -103.8437
24.4674 | -103.8513
24.471 | -103.8594
24.4747 | -103.868
24.4784 | -103.8768
24.482 | -103.8857
24.4857 | -103.8964
24.4893 | -103.9079
24.493 | -103.9203
24.4967 | -103.9332
24.5003 | -103.9452
24.504 | -103.9569
24.5077 | -103.9685
24.5113 | -103.9755
24.515 | -103.9857
24.5187 | -103.9975
24.5223 | -104.0067
24.526 | -104.0168
24.5297 | -104.0266
24.5333 | -104.0344
24.537 | -104.0418
24.5407 | -104.0489
24.5443 | -104.0564
24.548 | -104.0622
24.5517 | -104.0652
24.5553 | -104.0642
24.559 | -104.0606
24.5627 | -104.0548
24.5663 | -104.0469
24.57 | -104.0336
24.5737 | -104.0169
24.5773 | -103.9939
24.581 | -103.9538
24.5847 | -103.8765
24.5883 | -103.7832
24.592 | -103.6704
24.5956 | -103.5433
24.5993 | -103.4091
24.603 | -103.288
24.6066 | -103.1805
24.6103 | -103.0726
24.614 | -102.9639
24.6176 | -102.8483
24.6213 | -102.7248
24.625 | -102.5976
24.6286 | -102.4709
24.6323 | -102.3389
24.636 | -102.203
24.6396 | -102.0648
24.6433 | -101.9249
24.647 | -101.7818
24.6506 | -101.6375
24.6543 | -101.4923
24.658 | -101.345
24.6616 | -101.1975
24.6653 | -101.05
24.669 | -100.9028
24.6726 | -100.757
24.6763 | -100.6131
24.68 | -100.4701
24.6836 | -100.3289
24.6873 | -100.1906
24.691 | -100.0557
24.6946 | -99.925
24.6983 | -99.7982
24.7019 | -99.6764
24.7056 | -99.5591
24.7093 | -99.4459
24.7129 | -99.3384
24.7166 | -99.2363
24.7203 | -99.1403
24.7239 | -99.0496
24.7276 | -98.9648
24.7313 | -98.8868
24.7349 | -98.8153
24.7386 | -98.7508
24.7423 | -98.6931
24.7459 | -98.6421
24.7496 | -98.5976
24.7533 | -98.5605
24.7569 | -98.531
24.7606 | -98.5088
24.7643 | -98.493
24.7679 | -98.4844
24.7716 | -98.4831
24.7753 | -98.4893
24.7789 | -98.5023
24.7826 | -98.5222
24.7863 | -98.5499
24.7899 | -98.5853
24.7936 | -98.6273
24.7973 | -98.676
24.8009 | -98.7318
24.8046 | -98.794
24.8082 | -98.8632
24.8119 | -98.9385
24.8156 | -99.0198
24.8192 | -99.1069
24.8229 | -99.2003
24.8266 | -99.2993
24.8302 | -99.4029
24.8339 | -99.5114
24.8376 | -99.6241
24.8412 | -99.741
24.8449 | -99.8614
24.8486 | -99.9861
24.8522 | -100.1072
24.8559 | -100.2335
24.8596 | -100.3628
24.8632 | -100.4956
24.8669 | -100.6287
24.8706 | -100.7624
24.8742 | -100.8979
24.8779 | -101.0335
24.8816 | -101.1688
24.8852 | -101.3039
24.8889 | -101.4376
24.8926 | -101.5707
24.8962 | -101.7015
24.8999 | -101.8302
24.9036 | -101.9575
24.9072 | -102.0827
24.9109 | -102.2055
24.9145 | -102.3251
24.9182 | -102.4399
24.9219 | -102.5528
24.9255 | -102.6626
24.9292 | -102.7705
24.9329 | -102.8825
24.9365 | -103.0108
24.9402 | -103.1363
24.9439 | -103.2648
24.9475 | -103.3899
24.9512 | -103.5052
24.9549 | -103.5919
24.9585 | -103.6509
24.9622 | -103.6985
24.9659 | -103.7375
24.9695 | -103.7734
24.9732 | -103.8072
24.9769 | -103.8367
24.9805 | -103.8567
24.9842 | -103.8744
24.9879 | -103.889
24.9915 | -103.9011
24.9952 | -103.9099
24.9989 | -103.9195
25.0025 | -103.9272
25.0062 | -103.9341
25.0099 | -103.942
25.0135 | -103.9482
25.0172 | -103.9546
25.0208 | -103.9599
25.0245 | -103.9625
25.0282 | -103.9655
25.0318 | -103.9714
25.0355 | -103.9773
25.0392 | -103.9807
25.0428 | -103.9826
25.0465 | -103.9835
25.0502 | -103.9846
25.0538 | -103.984
25.0575 | -103.9847
25.0612 | -103.9846
25.0648 | -103.9824
25.0685 | -103.9808
25.0722 | -103.9773
25.0758 | -103.9755
25.0795 | -103.9725
25.0832 | -103.9688
25.0868 | -103.9657
25.0905 | -103.9632
25.0942 | -103.9606
25.0978 | -103.9562
25.1015 | -103.9522
25.1052 | -103.9484
25.1088 | -103.9451
25.1125 | -103.9428
25.1162 | -103.94
25.1198 | -103.9358
25.1235 | -103.9314
25.1271 | -103.9286
25.1308 | -103.9273
25.1345 | -103.9259
25.1381 | -103.9222
25.1418 | -103.9184
25.1455 | -103.9162
25.1491 | -103.9149
25.1528 | -103.9117
25.1565 | -103.9103
25.1601 | -103.9079
25.1638 | -103.9027
25.1675 | -103.8983
25.1711 | -103.8956
25.1748 | -103.8924
25.1785 | -103.8895
25.1821 | -103.8876
25.1858 | -103.8854
25.1895 | -103.8837
25.1931 | -103.8823
25.1968 | -103.8821
25.2005 | -103.8789
25.2041 | -103.8737
25.2078 | -103.8705
25.2115 | -103.867
25.2151 | -103.863
25.2188 | -103.8605
25.2225 | -103.8591
25.2261 | -103.8577
25.2298 | -103.8546
25.2334 | -103.8515
25.2371 | -103.8484
25.2408 | -103.846
25.2444 | -103.8436
25.2481 | -103.839
25.2518 | -103.8352
25.2554 | -103.8301
25.2591 | -103.8244
25.2628 | -103.8202
25.2664 | -103.8142
25.2701 | -103.8076
25.2738 | -103.8019
25.2774 | -103.7964
25.2811 | -103.7884
25.2848 | -103.7814
25.2884 | -103.7753
25.2921 | -103.7694
25.2958 | -103.7619
25.2994 | -103.7527
25.3031 | -103.742
25.3068 | -103.7323
25.3104 | -103.7206
25.3141 | -103.7104
25.3178 | -103.7009
25.3214 | -103.6909
25.3251 | -103.6818
25.3288 | -103.6722
25.3324 | -103.6614
25.3361 | -103.6531
25.3397 | -103.6447
25.3434 | -103.6354
25.3471 | -103.6251
25.3507 | -103.6159
25.3544 | -103.6082
25.3581 | -103.602
25.3617 | -103.5949
25.3654 | -103.5877
25.3691 | -103.5812
25.3727 | -103.575
25.3764 | -103.5688
25.3801 | -103.5652
25.3837 | -103.5616
25.3874 | -103.5558
25.3911 | -103.5499
25.3947 | -103.5455
25.3984 | -103.5411
25.4021 | -103.5367
25.4057 | -103.5325
25.4094 | -103.5291
25.4131 | -103.527
25.4167 | -103.5243
25.4204 | -103.5227
25.4241 | -103.5228
25.4277 | -103.5238
25.4314 | -103.5261
25.4351 | -103.5272
25.4387 | -103.5292
25.4424 | -103.531
25.446 | -103.5325
25.4497 | -103.5341
25.4534 | -103.5363
25.457 | -103.5401
25.4607 | -103.545
25.4644 | -103.5492
25.468 | -103.5543
25.4717 | -103.5595
25.4754 | -103.5663
25.479 | -103.5739
25.4827 | -103.5815
25.4864 | -103.5882
25.49 | -103.5948
25.4937 | -103.6021
25.4974 | -103.6082
25.501 | -103.6135
25.5047 | -103.619
25.5084 | -103.6243
25.512 | -103.6298
25.5157 | -103.6349
25.5194 | -103.6419
25.523 | -103.6501
25.5267 | -103.6587
25.5304 | -103.6654
25.534 | -103.672
25.5377 | -103.6787
25.5414 | -103.6848
25.545 | -103.6887
25.5487 | -103.6914
25.5523 | -103.6927
25.556 | -103.694
25.5597 | -103.6979
25.5633 | -103.7026
25.567 | -103.7052
25.5707 | -103.7087
25.5743 | -103.7105
25.578 | -103.7089
25.5817 | -103.7083
25.5853 | -103.7083
25.589 | -103.7102
25.5927 | -103.7107
25.5963 | -103.7106
25.6 | -103.7116
25.6037 | -103.7114
25.6073 | -103.7122
25.611 | -103.7135
25.6147 | -103.7137
25.6183 | -103.7133
25.622 | -103.7125
25.6257 | -103.7114
25.6293 | -103.7123
25.633 | -103.7124
25.6367 | -103.7097
25.6403 | -103.7089
25.644 | -103.7069
25.6477 | -103.7073
25.6513 | -103.7067
25.655 | -103.7064
25.6586 | -103.7065
25.6623 | -103.7074
25.666 | -103.7074
25.6696 | -103.7093
25.6733 | -103.7108
25.677 | -103.712
25.6806 | -103.7141
25.6843 | -103.7158
25.688 | -103.7179
25.6916 | -103.7223
25.6953 | -103.725
25.699 | -103.7287
25.7026 | -103.7323
25.7063 | -103.7364
25.71 | -103.74
25.7136 | -103.7433
25.7173 | -103.747
25.721 | -103.7521
25.7246 | -103.7571
25.7283 | -103.7635
25.732 | -103.7707
25.7356 | -103.7782
25.7393 | -103.7871
25.743 | -103.7952
25.7466 | -103.8026
25.7503 | -103.811
25.754 | -103.82
25.7576 | -103.8286
25.7613 | -103.8385
25.7649 | -103.8496
25.7686 | -103.8592
25.7723 | -103.8688
25.7759 | -103.88
25.7796 | -103.8907
25.7833 | -103.9021
25.7869 | -103.913
25.7906 | -103.9218
25.7943 | -103.9308
25.7979 | -103.9415
25.8016 | -103.9531
25.8053 | -103.9664
25.8089 | -103.9794
25.8126 | -103.9904
25.8163 | -103.9999
25.8199 | -104.0098
25.8236 | -104.0187
25.8273 | -104.0276
25.8309 | -104.0356
25.8346 | -104.0418
25.8383 | -104.0483
25.8419 | -104.0561
25.8456 | -104.0631
25.8493 | -104.0685
25.8529 | -104.0748
25.8566 | -104.0821
25.8603 | -104.0893
25.8639 | -104.0957
25.8676 | -104.1003
25.8712 | -104.1039
25.8749 | -104.1086
25.8786 | -104.1123
25.8822 | -104.1149
25.8859 | -104.1176
25.8896 | -104.1186
25.8932 | -104.119
25.8969 | -104.1183
25.9006 | -104.1156
25.9042 | -104.1142
25.9079 | -104.1143
25.9116 | -104.1109
25.9152 | -104.1102
25.9189 | -104.1079
25.9226 | -104.1033
25.9262 | -104.0991
25.9299 | -104.0946
25.9336 | -104.0912
25.9372 | -104.0862
25.9409 | -104.0824
25.9446 | -104.0805
25.9482 | -104.0767
25.9519 | -104.0729
25.9556 | -104.0675
25.9592 | -104.0635
25.9629 | -104.0593
25.9666 | -104.054
25.9702 | -104.0499
25.9739 | -104.0456
25.9775 | -104.0409
25.9812 | -104.0387
25.9849 | -104.037
25.9885 | -104.0351
25.9922 | -104.0349
25.9959 | -104.0337
25.9995 | -104.032
26.0032 | -104.0338
26.0069 | -104.0348
26.0105 | -104.0352
26.0142 | -104.0344
26.0179 | -104.0351
26.0215 | -104.038
26.0252 | -104.0402
26.0289 | -104.0424
26.0325 | -104.0448
26.0362 | -104.0476
26.0399 | -104.0494
26.0435 | -104.0512
26.0472 | -104.0531
26.0509 | -104.0566
26.0545 | -104.0615
26.0582 | -104.066
26.0619 | -104.0709
26.0655 | -104.0734
26.0692 | -104.0762
26.0729 | -104.081
26.0765 | -104.0862
26.0802 | -104.0909
26.0838 | -104.0958
26.0875 | -104.1002
26.0912 | -104.1058
26.0948 | -104.1094
26.0985 | -104.1114
26.1022 | -104.1148
26.1058 | -104.1183
26.1095 | -104.1223
26.1132 | -104.1258
26.1168 | -104.1297
26.1205 | -104.1321
26.1242 | -104.1331
26.1278 | -104.1345
26.1315 | -104.1366
26.1352 | -104.1386
26.1388 | -104.1403
26.1425 | -104.1402
26.1462 | -104.1396
26.1498 | -104.1391
26.1535 | -104.1386
26.1572 | -104.1384
26.1608 | -104.1374
26.1645 | -104.1374
26.1682 | -104.1362
26.1718 | -104.133
26.1755 | -104.1294
26.1792 | -104.1287
26.1828 | -104.1284
26.1865 | -104.1281
26.1901 | -104.127
26.1938 | -104.1256
26.1975 | -104.1238
26.2011 | -104.1223
26.2048 | -104.1211
26.2085 | -104.1197
26.2121 | -104.1185
26.2158 | -104.1174
26.2195 | -104.1153
26.2231 | -104.1115
26.2268 | -104.1093
26.2305 | -104.1069
26.2341 | -104.1049
26.2378 | -104.103
26.2415 | -104.0996
26.2451 | -104.0982
26.2488 | -104.0971
26.2525 | -104.0953
26.2561 | -104.0926
26.2598 | -104.0917
26.2635 | -104.0921
26.2671 | -104.0929
26.2708 | -104.0938
26.2745 | -104.0947
26.2781 | -104.0944
26.2818 | -104.0944
26.2855 | -104.0954
26.2891 | -104.0975
26.2928 | -104.0989
26.2964 | -104.0993
26.3001 | -104.1011
26.3038 | -104.1017
26.3074 | -104.1022
26.3111 | -104.1036
26.3148 | -104.1044
26.3184 | -104.1052
26.3221 | -104.107
26.3258 | -104.1107
26.3294 | -104.1123
26.3331 | -104.1126
26.3368 | -104.1133
26.3404 | -104.1145
26.3441 | -104.1159
26.3478 | -104.1178
26.3514 | -104.1212
26.3551 | -104.1253
26.3588 | -104.1284
26.3624 | -104.1289
26.3661 | -104.132
26.3698 | -104.1358
26.3734 | -104.1376
26.3771 | -104.139
26.3808 | -104.1426
26.3844 | -104.146
26.3881 | -104.1499
26.3918 | -104.1532
26.3954 | -104.1551
26.3991 | -104.1573
26.4027 | -104.1594
26.4064 | -104.1615
26.4101 | -104.163
26.4137 | -104.166
26.4174 | -104.1684
26.4211 | -104.1696
26.4247 | -104.1714
26.4284 | -104.1724
26.4321 | -104.1719
26.4357 | -104.1727
26.4394 | -104.1711
26.4431 | -104.172
26.4467 | -104.1721
26.4504 | -104.1715
26.4541 | -104.1702
26.4577 | -104.1684
26.4614 | -104.1673
26.4651 | -104.1641
26.4687 | -104.162
26.4724 | -104.1622
26.4761 | -104.1605
26.4797 | -104.1584
26.4834 | -104.1559
26.4871 | -104.1534
26.4907 | -104.1489
26.4944 | -104.144
26.4981 | -104.1411
26.5017 | -104.1388
26.5054 | -104.1351
26.509 | -104.1319
26.5127 | -104.1282
26.5164 | -104.1239
26.52 | -104.1218
26.5237 | -104.1195
26.5274 | -104.1155
26.531 | -104.1106
26.5347 | -104.1044
26.5384 | -104.096
26.542 | -104.0902
26.5457 | -104.086
26.5494 | -104.0825
26.553 | -104.0767
26.5567 | -104.0728
26.5604 | -104.0688
26.564 | -104.0647
26.5677 | -104.0599
26.5714 | -104.0553
26.575 | -104.0511
26.5787 | -104.0461
26.5824 | -104.04
26.586 | -104.0343
26.5897 | -104.0284
26.5934 | -104.0235
26.597 | -104.02
26.6007 | -104.0151
26.6044 | -104.0086
26.608 | -104.0053
26.6117 | -104.0025
26.6153 | -103.9997
26.619 | -103.9947
26.6227 | -103.9889
26.6263 | -103.9834
26.63 | -103.9797
26.6337 | -103.9753
26.6373 | -103.9706
26.641 | -103.966
26.6447 | -103.9619
26.6483 | -103.9584
26.652 | -103.9554
26.6557 | -103.9529
26.6593 | -103.9506
26.663 | -103.949
26.6667 | -103.9483
26.6703 | -103.9475
26.674 | -103.9452
26.6777 | -103.9447
26.6813 | -103.9445
26.685 | -103.9434
26.6887 | -103.9428
26.6923 | -103.9421
26.696 | -103.9405
26.6997 | -103.9397
26.7033 | -103.9412
26.707 | -103.9436
26.7107 | -103.9456
26.7143 | -103.947
26.718 | -103.9493
26.7216 | -103.9508
26.7253 | -103.9498
26.729 | -103.9497
26.7326 | -103.9505
26.7363 | -103.9535
26.74 | -103.9544
26.7436 | -103.9555
26.7473 | -103.957
26.751 | -103.9591
26.7546 | -103.9599
26.7583 | -103.9613
26.762 | -103.9607
26.7656 | -103.9617
26.7693 | -103.9628
26.773 | -103.9623
26.7766 | -103.9633
26.7803 | -103.9643
26.784 | -103.9666
26.7876 | -103.9678
26.7913 | -103.9687
26.795 | -103.9681
26.7986 | -103.9666
26.8023 | -103.9651
26.806 | -103.9649
26.8096 | -103.9648
26.8133 | -103.9635
26.817 | -103.9634
26.8206 | -103.9633
26.8243 | -103.9635
26.828 | -103.9649
26.8316 | -103.966
26.8353 | -103.9663
26.8389 | -103.9679
26.8426 | -103.9696
26.8463 | -103.9699
26.8499 | -103.9675
26.8536 | -103.9666
26.8573 | -103.9663
26.8609 | -103.9658
26.8646 | -103.9649
26.8683 | -103.9663
26.8719 | -103.9678
26.8756 | -103.9681
26.8793 | -103.9678
26.8829 | -103.9686
26.8866 | -103.9699
26.8903 | -103.9707
26.8939 | -103.9733
26.8976 | -103.9759
26.9013 | -103.9782
26.9049 | -103.9803
26.9086 | -103.9827
26.9123 | -103.987
26.9159 | -103.9917
26.9196 | -103.9946
26.9233 | -103.9977
26.9269 | -103.9998
26.9306 | -104.0033
26.9342 | -104.0064
26.9379 | -104.0089
26.9416 | -104.0122
26.9452 | -104.0164
26.9489 | -104.0212
26.9526 | -104.0264
26.9562 | -104.0322
26.9599 | -104.038
26.9636 | -104.0431
26.9672 | -104.0514
26.9709 | -104.0581
26.9746 | -104.0641
26.9782 | -104.0675
26.9819 | -104.0705
26.9856 | -104.0757
26.9892 | -104.0803
26.9929 | -104.0847
26.9966 | -104.0882
27.0002 | -104.0917
27.0039 | -104.0972
27.0076 | -104.1043
27.0112 | -104.11
27.0149 | -104.1146
27.0186 | -104.118
27.0222 | -104.123
27.0259 | -104.127
27.0296 | -104.1299
27.0332 | -104.1345
27.0369 | -104.1392
27.0405 | -104.1426
27.0442 | -104.1468
27.0479 | -104.1503
27.0515 | -104.1521
27.0552 | -104.1544
27.0589 | -104.1566
27.0625 | -104.1587
27.0662 | -104.1603
27.0699 | -104.1616
27.0735 | -104.1627
27.0772 | -104.1643
27.0809 | -104.1642
27.0845 | -104.1632
27.0882 | -104.1631
27.0919 | -104.1618
27.0955 | -104.1603
27.0992 | -104.1578
27.1029 | -104.1567
27.1065 | -104.1566
27.1102 | -104.1556
27.1139 | -104.1526
27.1175 | -104.1478
27.1212 | -104.1427
27.1249 | -104.1389
27.1285 | -104.1359
27.1322 | -104.1323
27.1359 | -104.1288
27.1395 | -104.1248
27.1432 | -104.1227
27.1469 | -104.1185
27.1505 | -104.1158
27.1542 | -104.1139
27.1578 | -104.1114
27.1615 | -104.1086
27.1652 | -104.1053
27.1688 | -104.1017
27.1725 | -104.0961
27.1762 | -104.0937
27.1798 | -104.0918
27.1835 | -104.0896
27.1872 | -104.0866
27.1908 | -104.0828
27.1945 | -104.0796
27.1982 | -104.0769
27.2018 | -104.0741
27.2055 | -104.0719
27.2092 | -104.0686
27.2128 | -104.0664
27.2165 | -104.0633
27.2202 | -104.0593
27.2238 | -104.0577
27.2275 | -104.057
27.2312 | -104.0562
27.2348 | -104.0543
27.2385 | -104.0524
27.2422 | -104.0506
27.2458 | -104.0505
27.2495 | -104.0483
27.2532 | -104.0455
27.2568 | -104.0431
27.2605 | -104.0393
27.2641 | -104.0347
27.2678 | -104.0313
27.2715 | -104.0295
27.2751 | -104.0259
27.2788 | -104.0205
27.2825 | -104.0151
27.2861 | -104.0122
27.2898 | -104.0107
27.2935 | -104.0065
27.2971 | -104.0029
27.3008 | -103.9993
27.3045 | -103.9944
27.3081 | -103.9906
27.3118 | -103.986
27.3155 | -103.9816
27.3191 | -103.9779
27.3228 | -103.9741
27.3265 | -103.9702
27.3301 | -103.9662
27.3338 | -103.9626
27.3375 | -103.9594
27.3411 | -103.9577
27.3448 | -103.955
27.3485 | -103.9513
27.3521 | -103.9469
27.3558 | -103.9439
27.3594 | -103.9426
27.3631 | -103.9432
27.3668 | -103.9441
27.3704 | -103.9462
27.3741 | -103.9464
27.3778 | -103.9468
27.3814 | -103.9479
27.3851 | -103.9493
27.3888 | -103.9497
27.3924 | -103.9498
27.3961 | -103.9507
27.3998 | -103.9522
27.4034 | -103.9543
27.4071 | -103.9561
27.4108 | -103.9596
27.4144 | -103.9643
27.4181 | -103.9671
27.4218 | -103.9702
27.4254 | -103.975
27.4291 | -103.9794
27.4328 | -103.9834
27.4364 | -103.9896
27.4401 | -103.9967
27.4438 | -104.0012
27.4474 | -104.006
27.4511 | -104.0127
27.4548 | -104.0195
27.4584 | -104.0256
27.4621 | -104.0315
27.4657 | -104.036
27.4694 | -104.0396
27.4731 | -104.0436
27.4767 | -104.0482
27.4804 | -104.0525
27.4841 | -104.0551
27.4877 | -104.0578
27.4914 | -104.0617
27.4951 | -104.064
27.4987 | -104.0661
27.5024 | -104.07
27.5061 | -104.076
27.5097 | -104.081
27.5134 | -104.0842
27.5171 | -104.0862
27.5207 | -104.0876
27.5244 | -104.0891
27.5281 | -104.0899
27.5317 | -104.0894
27.5354 | -104.0901
27.5391 | -104.0892
27.5427 | -104.0869
27.5464 | -104.0839
27.5501 | -104.0815
27.5537 | -104.0794
27.5574 | -104.0775
27.5611 | -104.077
27.5647 | -104.0747
27.5684 | -104.072
27.5721 | -104.0688
27.5757 | -104.0652
27.5794 | -104.0621
27.583 | -104.0588
27.5867 | -104.0543
27.5904 | -104.0496
27.594 | -104.0438
27.5977 | -104.0396
27.6014 | -104.0356
27.605 | -104.0315
27.6087 | -104.0253
27.6124 | -104.018
27.616 | -104.0116
27.6197 | -104.0071
27.6234 | -104.003
27.627 | -103.9965
27.6307 | -103.9884
27.6344 | -103.98
27.638 | -103.9735
27.6417 | -103.9678
27.6454 | -103.9621
27.649 | -103.9577
27.6527 | -103.9553
27.6564 | -103.9514
27.66 | -103.9453
27.6637 | -103.939
27.6674 | -103.9342
27.671 | -103.9304
27.6747 | -103.9262
27.6784 | -103.9222
27.682 | -103.9211
27.6857 | -103.918
27.6893 | -103.9144
27.693 | -103.9127
27.6967 | -103.9111
27.7003 | -103.9108
27.704 | -103.9087
27.7077 | -103.9062
27.7113 | -103.9032
27.715 | -103.9019
27.7187 | -103.902
27.7223 | -103.9019
27.726 | -103.9025
27.7297 | -103.903
27.7333 | -103.9044
27.737 | -103.9063
27.7407 | -103.9093
27.7443 | -103.9112
27.748 | -103.9127
27.7517 | -103.9152
27.7553 | -103.9177
27.759 | -103.9186
27.7627 | -103.9199
27.7663 | -103.9205
27.77 | -103.9229
27.7737 | -103.9259
27.7773 | -103.9278
27.781 | -103.9298
27.7847 | -103.933
27.7883 | -103.9371
27.792 | -103.9413
27.7956 | -103.9445
27.7993 | -103.9447
27.803 | -103.9439
27.8066 | -103.9438
27.8103 | -103.9445
27.814 | -103.946
27.8176 | -103.9456
27.8213 | -103.9465
27.825 | -103.9465
27.8286 | -103.9466
27.8323 | -103.9484
27.836 | -103.9503
27.8396 | -103.9513
27.8433 | -103.9506
27.847 | -103.9499
27.8506 | -103.9472
27.8543 | -103.947
27.858 | -103.9479
27.8616 | -103.9462
27.8653 | -103.9443
27.869 | -103.9432
27.8726 | -103.9432
27.8763 | -103.9409
27.88 | -103.9374
27.8836 | -103.933
27.8873 | -103.9284
27.891 | -103.9246
27.8946 | -103.9191
27.8983 | -103.912
27.9019 | -103.9063
27.9056 | -103.9005
27.9093 | -103.8956
27.9129 | -103.8921
27.9166 | -103.8859
27.9203 | -103.8792
27.9239 | -103.8745
27.9276 | -103.8688
27.9313 | -103.8632
27.9349 | -103.858
27.9386 | -103.8515
27.9423 | -103.8454
27.9459 | -103.8397
27.9496 | -103.8344
27.9533 | -103.8284
27.9569 | -103.8229
27.9606 | -103.8161
27.9643 | -103.8111
27.9679 | -103.8058
27.9716 | -103.8001
27.9753 | -103.7952
27.9789 | -103.7904
27.9826 | -103.7857
27.9863 | -103.7813
27.9899 | -103.7775
27.9936 | -103.7733
27.9973 | -103.7712
28.0009 | -103.7713