## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 219 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 218
Hood | 02-92B/P | 156
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1011.7 | hPa |  | PRC PRO 01
Humidity | 36.3 | %/Hr
Internal temperature of calibration bench | 23.97 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1014 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.3 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.06 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.99 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3108 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 03/05/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 14/05/2018
Engineer | Anthony HUE | Approval |  | 18/05/2018
CEO | Robert GLEMOT | Approval |  | 18/05/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4601 | 57.5852
0.012209 | -36.0303 | 38.4375
0.018313 | -34.9735 | 28.2789
0.024417 | -34.5108 | 22.0256
0.030521 | -34.3933 | 18.2345
0.036626 | -34.2475 | 15.172
0.04273 | -34.1521 | 12.8105
0.048834 | -34.1064 | 11.2803
0.054938 | -34.044 | 10.1303
0.061043 | -33.9837 | 9.0515
0.067147 | -33.9721 | 8.2278
0.073251 | -33.9719 | 7.4644
0.079355 | -33.9885 | 6.8025
0.08546 | -34.0011 | 6.2148
0.091564 | -33.9959 | 5.7778
0.097668 | -33.9788 | 5.5218
0.10377 | -33.9583 | 5.2666
0.10988 | -33.9421 | 4.9856
0.11598 | -33.9333 | 4.6927
0.12209 | -33.9345 | 4.4099
0.12819 | -33.9433 | 4.1275
0.13429 | -33.9531 | 3.8408
0.1404 | -33.9562 | 3.5682
0.1465 | -33.958 | 3.3158
0.15261 | -33.9556 | 3.1012
0.15871 | -33.9496 | 2.9428
0.16482 | -33.9419 | 2.8404
0.17092 | -33.9395 | 2.7956
0.17702 | -33.9373 | 2.7197
0.18313 | -33.9362 | 2.6375
0.18923 | -33.9355 | 2.5257
0.19534 | -33.9341 | 2.3974
0.20144 | -33.9328 | 2.2549
0.20754 | -33.9301 | 2.135
0.21365 | -33.9292 | 2.0211
0.21975 | -33.9287 | 1.924
0.22586 | -33.9298 | 1.8381
0.23196 | -33.9321 | 1.767
0.23807 | -33.936 | 1.7051
0.24417 | -33.94 | 1.6451
0.25027 | -33.944 | 1.5817
0.25638 | -33.9465 | 1.518
0.26248 | -33.9484 | 1.4495
0.26859 | -33.9495 | 1.3823
0.27469 | -33.95 | 1.31
0.2808 | -33.951 | 1.2375
0.2869 | -33.9517 | 1.1675
0.293 | -33.9526 | 1.1017
0.29911 | -33.9534 | 1.0433
0.30521 | -33.9543 | 0.98628
0.31132 | -33.955 | 0.94918
0.31742 | -33.9562 | 0.9076
0.32353 | -33.9565 | 0.87355
0.32963 | -33.9575 | 0.84484
0.33573 | -33.9586 | 0.82392
0.34184 | -33.96 | 0.80161
0.34794 | -33.9622 | 0.77634
0.35405 | -33.9644 | 0.75104
0.36015 | -33.9668 | 0.72451
0.36626 | -33.9688 | 0.69589
0.37236 | -33.9709 | 0.65857
0.37846 | -33.973 | 0.61426
0.38457 | -33.9738 | 0.55845
0.39067 | -33.9748 | 0.51225
0.39678 | -33.9754 | 0.46155
0.40288 | -33.9757 | 0.40896
0.40899 | -33.9756 | 0.35859
0.41509 | -33.9749 | 0.30945
0.42119 | -33.9748 | 0.26765
0.4273 | -33.9747 | 0.22281
0.4334 | -33.9746 | 0.17817
0.43951 | -33.9748 | 0.13604
0.44561 | -33.9756 | 0.097358
0.45172 | -33.977 | 0.066048
0.45782 | -33.9777 | 0.029766
0.46392 | -33.9789 | -0.0058708
0.47003 | -33.9799 | -0.037802
0.47613 | -33.9807 | -0.071466
0.48224 | -33.9812 | -0.10568
0.48834 | -33.9803 | -0.14009
0.49445 | -33.9801 | -0.17478
0.50055 | -33.9796 | -0.20655
0.50665 | -33.9789 | -0.24042
0.51276 | -33.978 | -0.27419
0.51886 | -33.977 | -0.30778
0.52497 | -33.9759 | -0.3321
0.53107 | -33.9751 | -0.36154
0.53717 | -33.9745 | -0.38969
0.54328 | -33.9737 | -0.41278
0.54938 | -33.9735 | -0.4356
0.55549 | -33.9735 | -0.45681
0.56159 | -33.9746 | -0.46921
0.5677 | -33.9749 | -0.4845
0.5738 | -33.9759 | -0.50137
0.5799 | -33.9766 | -0.51701
0.58601 | -33.9777 | -0.52812
0.59211 | -33.9793 | -0.53545
0.59822 | -33.9806 | -0.54577
0.60432 | -33.982 | -0.55859
0.61043 | -33.9833 | -0.56748
0.61653 | -33.9844 | -0.57852
0.62263 | -33.986 | -0.58868
0.62874 | -33.9866 | -0.59997
0.63484 | -33.9875 | -0.60998
0.64095 | -33.9879 | -0.62081
0.64705 | -33.9883 | -0.63127
0.65316 | -33.9886 | -0.64426
0.65926 | -33.9887 | -0.66037
0.66536 | -33.9876 | -0.67791
0.67147 | -33.9873 | -0.69156
0.67757 | -33.987 | -0.70616
0.68368 | -33.9863 | -0.72512
0.68978 | -33.9856 | -0.74251
0.69589 | -33.9847 | -0.76153
0.70199 | -33.9839 | -0.78404
0.70809 | -33.9831 | -0.80319
0.7142 | -33.9824 | -0.82357
0.7203 | -33.9814 | -0.84278
0.72641 | -33.9808 | -0.86504
0.73251 | -33.9802 | -0.88496
0.73862 | -33.9805 | -0.90648
0.74472 | -33.9802 | -0.92634
0.75082 | -33.98 | -0.94822
0.75693 | -33.9803 | -0.96984
0.76303 | -33.9809 | -0.99158
0.76914 | -33.9827 | -1.0118
0.77524 | -33.9836 | -1.0313
0.78135 | -33.9848 | -1.0513
0.78745 | -33.9862 | -1.071
0.79355 | -33.9874 | -1.0906
0.79966 | -33.9891 | -1.1099
0.80576 | -33.9912 | -1.1294
0.81187 | -33.9928 | -1.1507
0.81797 | -33.9947 | -1.1717
0.82408 | -33.9963 | -1.192
0.83018 | -33.998 | -1.2123
0.83628 | -33.9998 | -1.2307
0.84239 | -34.0009 | -1.2511
0.84849 | -34.0023 | -1.2725
0.8546 | -34.0037 | -1.2921
0.8607 | -34.0049 | -1.3126
0.86681 | -34.006 | -1.3315
0.87291 | -34.0069 | -1.3502
0.87901 | -34.0074 | -1.3682
0.88512 | -34.0081 | -1.3863
0.89122 | -34.0089 | -1.4045
0.89733 | -34.0093 | -1.4202
0.90343 | -34.0097 | -1.4378
0.90953 | -34.0099 | -1.4523
0.91564 | -34.01 | -1.4682
0.92174 | -34.0102 | -1.4833
0.92785 | -34.0103 | -1.4981
0.93395 | -34.0104 | -1.5122
0.94006 | -34.0106 | -1.5252
0.94616 | -34.0104 | -1.5373
0.95226 | -34.0105 | -1.5502
0.95837 | -34.0106 | -1.5643
0.96447 | -34.0108 | -1.5773
0.97058 | -34.0107 | -1.5908
0.97668 | -34.0105 | -1.6043
0.98279 | -34.0103 | -1.6173
0.98889 | -34.0101 | -1.6282
0.99499 | -34.0101 | -1.6418
1.0011 | -34.01 | -1.6528
1.0072 | -34.0097 | -1.6646
1.0133 | -34.0094 | -1.6757
1.0194 | -34.009 | -1.6887
1.0255 | -34.0087 | -1.7026
1.0316 | -34.0082 | -1.7147
1.0377 | -34.0078 | -1.7284
1.0438 | -34.0073 | -1.7413
1.0499 | -34.0066 | -1.7565
1.056 | -34.006 | -1.7708
1.0621 | -34.0054 | -1.7851
1.0682 | -34.005 | -1.7977
1.0743 | -34.0044 | -1.8131
1.0805 | -34.0039 | -1.8274
1.0866 | -34.003 | -1.8425
1.0927 | -34.0022 | -1.8569
1.0988 | -34.0017 | -1.8718
1.1049 | -34.0009 | -1.8859
1.111 | -34.0002 | -1.9012
1.1171 | -33.9995 | -1.9157
1.1232 | -33.9991 | -1.9308
1.1293 | -33.9987 | -1.9447
1.1354 | -33.9984 | -1.9596
1.1415 | -33.9981 | -1.975
1.1476 | -33.9975 | -1.9904
1.1537 | -33.997 | -2.005
1.1598 | -33.9971 | -2.0197
1.1659 | -33.9969 | -2.0329
1.172 | -33.9966 | -2.0468
1.1781 | -33.9964 | -2.0607
1.1842 | -33.9963 | -2.0748
1.1903 | -33.9966 | -2.0859
1.1964 | -33.9965 | -2.0993
1.2025 | -33.9966 | -2.1132
1.2086 | -33.9969 | -2.1256
1.2147 | -33.9969 | -2.1381
1.2209 | -33.9973 | -2.1494
1.227 | -33.998 | -2.1602
1.2331 | -33.9983 | -2.1722
1.2392 | -33.9987 | -2.1842
1.2453 | -33.9991 | -2.1953
1.2514 | -33.9995 | -2.2054
1.2575 | -34.0001 | -2.2164
1.2636 | -34.0009 | -2.2256
1.2697 | -34.0014 | -2.235
1.2758 | -34.0021 | -2.2462
1.2819 | -34.0027 | -2.2563
1.288 | -34.0034 | -2.2672
1.2941 | -34.0042 | -2.2759
1.3002 | -34.0048 | -2.2843
1.3063 | -34.0053 | -2.2935
1.3124 | -34.006 | -2.3027
1.3185 | -34.0067 | -2.312
1.3246 | -34.0073 | -2.3213
1.3307 | -34.0082 | -2.3301
1.3368 | -34.0089 | -2.339
1.3429 | -34.0095 | -2.3482
1.349 | -34.01 | -2.3577
1.3551 | -34.0107 | -2.3673
1.3613 | -34.0112 | -2.3758
1.3674 | -34.0118 | -2.3859
1.3735 | -34.0123 | -2.3958
1.3796 | -34.0127 | -2.4051
1.3857 | -34.013 | -2.4145
1.3918 | -34.0135 | -2.4248
1.3979 | -34.0141 | -2.4341
1.404 | -34.0146 | -2.4454
1.4101 | -34.015 | -2.4559
1.4162 | -34.0153 | -2.4666
1.4223 | -34.0158 | -2.478
1.4284 | -34.0163 | -2.488
1.4345 | -34.0165 | -2.4996
1.4406 | -34.0168 | -2.5129
1.4467 | -34.0173 | -2.5245
1.4528 | -34.0177 | -2.5358
1.4589 | -34.0182 | -2.5477
1.465 | -34.0184 | -2.5585
1.4711 | -34.0189 | -2.5699
1.4772 | -34.0194 | -2.5831
1.4833 | -34.0198 | -2.596
1.4894 | -34.0203 | -2.6081
1.4955 | -34.0209 | -2.62
1.5016 | -34.0214 | -2.6325
1.5078 | -34.0218 | -2.6462
1.5139 | -34.0221 | -2.6591
1.52 | -34.0226 | -2.6716
1.5261 | -34.0229 | -2.6841
1.5322 | -34.0232 | -2.698
1.5383 | -34.0236 | -2.7104
1.5444 | -34.0241 | -2.723
1.5505 | -34.0245 | -2.7362
1.5566 | -34.025 | -2.7491
1.5627 | -34.0255 | -2.7627
1.5688 | -34.0261 | -2.7761
1.5749 | -34.0264 | -2.7887
1.581 | -34.0267 | -2.8019
1.5871 | -34.0271 | -2.8154
1.5932 | -34.0275 | -2.8282
1.5993 | -34.0279 | -2.8422
1.6054 | -34.0282 | -2.8553
1.6115 | -34.0287 | -2.8692
1.6176 | -34.0289 | -2.8817
1.6237 | -34.0291 | -2.8952
1.6298 | -34.0294 | -2.9076
1.6359 | -34.0298 | -2.9202
1.642 | -34.0302 | -2.9321
1.6482 | -34.0304 | -2.9449
1.6543 | -34.0307 | -2.9584
1.6604 | -34.031 | -2.9703
1.6665 | -34.0312 | -2.9827
1.6726 | -34.0314 | -2.9951
1.6787 | -34.0317 | -3.0075
1.6848 | -34.0316 | -3.0199
1.6909 | -34.0319 | -3.0325
1.697 | -34.0321 | -3.0446
1.7031 | -34.0322 | -3.057
1.7092 | -34.0321 | -3.069
1.7153 | -34.0323 | -3.0809
1.7214 | -34.0321 | -3.093
1.7275 | -34.0321 | -3.1055
1.7336 | -34.0321 | -3.1183
1.7397 | -34.0322 | -3.1305
1.7458 | -34.0322 | -3.1426
1.7519 | -34.0322 | -3.1546
1.758 | -34.0323 | -3.1664
1.7641 | -34.0324 | -3.178
1.7702 | -34.0324 | -3.1907
1.7763 | -34.0324 | -3.2021
1.7824 | -34.0323 | -3.2144
1.7885 | -34.0323 | -3.2267
1.7947 | -34.0325 | -3.2386
1.8008 | -34.0325 | -3.2498
1.8069 | -34.0325 | -3.2613
1.813 | -34.0325 | -3.2733
1.8191 | -34.0325 | -3.2859
1.8252 | -34.0326 | -3.2985
1.8313 | -34.0327 | -3.3107
1.8374 | -34.0329 | -3.3225
1.8435 | -34.0329 | -3.3343
1.8496 | -34.033 | -3.3472
1.8557 | -34.0332 | -3.3599
1.8618 | -34.0334 | -3.372
1.8679 | -34.0335 | -3.384
1.874 | -34.0336 | -3.3959
1.8801 | -34.0338 | -3.4083
1.8862 | -34.034 | -3.4201
1.8923 | -34.0341 | -3.4323
1.8984 | -34.0342 | -3.4443
1.9045 | -34.0343 | -3.4559
1.9106 | -34.0344 | -3.4681
1.9167 | -34.0344 | -3.4802
1.9228 | -34.0345 | -3.4921
1.9289 | -34.0346 | -3.504
1.9351 | -34.0348 | -3.5149
1.9412 | -34.035 | -3.5261
1.9473 | -34.0353 | -3.5382
1.9534 | -34.0355 | -3.5497
1.9595 | -34.0357 | -3.5615
1.9656 | -34.0359 | -3.5738
1.9717 | -34.0361 | -3.5855
1.9778 | -34.0364 | -3.597
1.9839 | -34.0367 | -3.6085
1.99 | -34.0368 | -3.6196
1.9961 | -34.0369 | -3.6318
2.0022 | -34.0372 | -3.6433
2.0083 | -34.0373 | -3.6546
2.0144 | -34.0375 | -3.666
2.0205 | -34.0378 | -3.6778
2.0266 | -34.0381 | -3.6892
2.0327 | -34.0383 | -3.7009
2.0388 | -34.0387 | -3.7125
2.0449 | -34.0389 | -3.724
2.051 | -34.0392 | -3.7354
2.0571 | -34.0395 | -3.7468
2.0632 | -34.0397 | -3.7584
2.0693 | -34.0398 | -3.7706
2.0754 | -34.0401 | -3.782
2.0816 | -34.0404 | -3.7942
2.0877 | -34.0406 | -3.8057
2.0938 | -34.041 | -3.8179
2.0999 | -34.0411 | -3.8296
2.106 | -34.0413 | -3.8413
2.1121 | -34.0415 | -3.8533
2.1182 | -34.0418 | -3.8641
2.1243 | -34.0421 | -3.875
2.1304 | -34.0423 | -3.8862
2.1365 | -34.0426 | -3.898
2.1426 | -34.0428 | -3.91
2.1487 | -34.0431 | -3.9213
2.1548 | -34.0435 | -3.9331
2.1609 | -34.0439 | -3.9446
2.167 | -34.0443 | -3.9559
2.1731 | -34.0446 | -3.9679
2.1792 | -34.0448 | -3.979
2.1853 | -34.0452 | -3.991
2.1914 | -34.0455 | -4.0024
2.1975 | -34.0458 | -4.014
2.2036 | -34.0461 | -4.0261
2.2097 | -34.0463 | -4.0376
2.2158 | -34.0467 | -4.0483
2.222 | -34.047 | -4.0589
2.2281 | -34.0472 | -4.0704
2.2342 | -34.0475 | -4.0818
2.2403 | -34.0477 | -4.0925
2.2464 | -34.048 | -4.1032
2.2525 | -34.0482 | -4.1147
2.2586 | -34.0486 | -4.1253
2.2647 | -34.0489 | -4.1358
2.2708 | -34.0492 | -4.1471
2.2769 | -34.0495 | -4.1579
2.283 | -34.0497 | -4.169
2.2891 | -34.05 | -4.18
2.2952 | -34.0501 | -4.1906
2.3013 | -34.0504 | -4.2015
2.3074 | -34.0506 | -4.2121
2.3135 | -34.0509 | -4.2235
2.3196 | -34.0511 | -4.2337
2.3257 | -34.0513 | -4.2445
2.3318 | -34.0514 | -4.2557
2.3379 | -34.0516 | -4.2665
2.344 | -34.0518 | -4.2779
2.3501 | -34.052 | -4.2886
2.3562 | -34.0522 | -4.299
2.3623 | -34.0523 | -4.3095
2.3685 | -34.0524 | -4.3201
2.3746 | -34.0526 | -4.3313
2.3807 | -34.0528 | -4.3419
2.3868 | -34.053 | -4.3526
2.3929 | -34.0532 | -4.3634
2.399 | -34.0532 | -4.3739
2.4051 | -34.0535 | -4.3845
2.4112 | -34.0537 | -4.3952
2.4173 | -34.0539 | -4.4058
2.4234 | -34.054 | -4.4167
2.4295 | -34.0541 | -4.4275
2.4356 | -34.0543 | -4.4384
2.4417 | -34.0544 | -4.4496
2.4478 | -34.0546 | -4.4605
2.4539 | -34.0547 | -4.4708
2.46 | -34.0548 | -4.4822
2.4661 | -34.055 | -4.4933
2.4722 | -34.0552 | -4.5036
2.4783 | -34.0553 | -4.514
2.4844 | -34.0553 | -4.5251
2.4905 | -34.0554 | -4.5361
2.4966 | -34.0556 | -4.5469
2.5027 | -34.0557 | -4.5574
2.5089 | -34.0558 | -4.5683
2.515 | -34.0559 | -4.5792
2.5211 | -34.056 | -4.5899
2.5272 | -34.056 | -4.6008
2.5333 | -34.0562 | -4.6118
2.5394 | -34.0563 | -4.6226
2.5455 | -34.0565 | -4.634
2.5516 | -34.0566 | -4.6444
2.5577 | -34.0568 | -4.6558
2.5638 | -34.0569 | -4.6675
2.5699 | -34.057 | -4.6789
2.576 | -34.057 | -4.6904
2.5821 | -34.0572 | -4.7017
2.5882 | -34.0573 | -4.7128
2.5943 | -34.0574 | -4.7236
2.6004 | -34.0576 | -4.7346
2.6065 | -34.0577 | -4.745
2.6126 | -34.0578 | -4.7555
2.6187 | -34.0579 | -4.766
2.6248 | -34.0581 | -4.7768
2.6309 | -34.0582 | -4.7881
2.637 | -34.0585 | -4.7997
2.6431 | -34.0586 | -4.8102
2.6492 | -34.0587 | -4.8215
2.6554 | -34.0589 | -4.8324
2.6615 | -34.059 | -4.8433
2.6676 | -34.0593 | -4.8547
2.6737 | -34.0595 | -4.8659
2.6798 | -34.0598 | -4.8769
2.6859 | -34.06 | -4.8882
2.692 | -34.0601 | -4.899
2.6981 | -34.0603 | -4.9102
2.7042 | -34.0605 | -4.9213
2.7103 | -34.0607 | -4.9319
2.7164 | -34.061 | -4.9429
2.7225 | -34.0612 | -4.9542
2.7286 | -34.0614 | -4.9652
2.7347 | -34.0616 | -4.9757
2.7408 | -34.0619 | -4.9859
2.7469 | -34.0621 | -4.9971
2.753 | -34.0624 | -5.0077
2.7591 | -34.0627 | -5.019
2.7652 | -34.0629 | -5.0304
2.7713 | -34.063 | -5.0411
2.7774 | -34.0633 | -5.0521
2.7835 | -34.0634 | -5.063
2.7896 | -34.0636 | -5.0735
2.7958 | -34.0639 | -5.0843
2.8019 | -34.0642 | -5.0947
2.808 | -34.0644 | -5.1054
2.8141 | -34.0647 | -5.1166
2.8202 | -34.065 | -5.1274
2.8263 | -34.0653 | -5.1383
2.8324 | -34.0657 | -5.1496
2.8385 | -34.0661 | -5.1604
2.8446 | -34.0664 | -5.1711
2.8507 | -34.0668 | -5.182
2.8568 | -34.0671 | -5.1924
2.8629 | -34.0673 | -5.2031
2.869 | -34.0676 | -5.2141
2.8751 | -34.068 | -5.2248
2.8812 | -34.0684 | -5.2361
2.8873 | -34.0686 | -5.247
2.8934 | -34.069 | -5.258
2.8995 | -34.0693 | -5.2691
2.9056 | -34.0695 | -5.2802
2.9117 | -34.0699 | -5.2912
2.9178 | -34.0702 | -5.3019
2.9239 | -34.0706 | -5.3134
2.93 | -34.071 | -5.3248
2.9361 | -34.0713 | -5.3356
2.9423 | -34.0716 | -5.3463
2.9484 | -34.072 | -5.3573
2.9545 | -34.0724 | -5.3683
2.9606 | -34.0728 | -5.3796
2.9667 | -34.0731 | -5.3908
2.9728 | -34.0735 | -5.402
2.9789 | -34.0739 | -5.4131
2.985 | -34.0743 | -5.4244
2.9911 | -34.0747 | -5.4357
2.9972 | -34.0751 | -5.4467
3.0033 | -34.0755 | -5.4582
3.0094 | -34.0759 | -5.4692
3.0155 | -34.0763 | -5.4802
3.0216 | -34.0767 | -5.4914
3.0277 | -34.077 | -5.503
3.0338 | -34.0774 | -5.5144
3.0399 | -34.0778 | -5.5257
3.046 | -34.0781 | -5.5372
3.0521 | -34.0784 | -5.5483
3.0582 | -34.0787 | -5.5603
3.0643 | -34.079 | -5.5718
3.0704 | -34.0794 | -5.5835
3.0765 | -34.0796 | -5.5946
3.0827 | -34.08 | -5.6059
3.0888 | -34.0803 | -5.6171
3.0949 | -34.0807 | -5.6285
3.101 | -34.081 | -5.6399
3.1071 | -34.0813 | -5.6516
3.1132 | -34.0817 | -5.6634
3.1193 | -34.0821 | -5.6748
3.1254 | -34.0825 | -5.6861
3.1315 | -34.0828 | -5.698
3.1376 | -34.0831 | -5.7094
3.1437 | -34.0834 | -5.721
3.1498 | -34.0838 | -5.733
3.1559 | -34.0841 | -5.7444
3.162 | -34.0844 | -5.7565
3.1681 | -34.0846 | -5.7684
3.1742 | -34.0849 | -5.7799
3.1803 | -34.0853 | -5.7915
3.1864 | -34.0856 | -5.8028
3.1925 | -34.0859 | -5.8144
3.1986 | -34.0862 | -5.8263
3.2047 | -34.0865 | -5.8383
3.2108 | -34.0867 | -5.8495
3.2169 | -34.0871 | -5.8612
3.223 | -34.0873 | -5.8729
3.2292 | -34.0875 | -5.8847
3.2353 | -34.0877 | -5.8956
3.2414 | -34.088 | -5.907
3.2475 | -34.0882 | -5.9184
3.2536 | -34.0885 | -5.9299
3.2597 | -34.0888 | -5.9408
3.2658 | -34.089 | -5.9518
3.2719 | -34.0892 | -5.9629
3.278 | -34.0895 | -5.9744
3.2841 | -34.0897 | -5.9856
3.2902 | -34.09 | -5.997
3.2963 | -34.0902 | -6.0085
3.3024 | -34.0903 | -6.0196
3.3085 | -34.0906 | -6.0312
3.3146 | -34.0908 | -6.0423
3.3207 | -34.091 | -6.0538
3.3268 | -34.0913 | -6.0653
3.3329 | -34.0915 | -6.0761
3.339 | -34.0916 | -6.0869
3.3451 | -34.0918 | -6.0982
3.3512 | -34.092 | -6.1089
3.3573 | -34.0922 | -6.1202
3.3634 | -34.0924 | -6.1313
3.3696 | -34.0926 | -6.1427
3.3757 | -34.0928 | -6.1535
3.3818 | -34.093 | -6.1644
3.3879 | -34.0932 | -6.1755
3.394 | -34.0933 | -6.1869
3.4001 | -34.0935 | -6.1977
3.4062 | -34.0936 | -6.2087
3.4123 | -34.0938 | -6.2194
3.4184 | -34.094 | -6.2306
3.4245 | -34.0942 | -6.2415
3.4306 | -34.0943 | -6.2522
3.4367 | -34.0945 | -6.2628
3.4428 | -34.0947 | -6.2732
3.4489 | -34.0948 | -6.2842
3.455 | -34.0949 | -6.2951
3.4611 | -34.0951 | -6.3063
3.4672 | -34.0953 | -6.3174
3.4733 | -34.0955 | -6.3285
3.4794 | -34.0956 | -6.3393
3.4855 | -34.0958 | -6.3498
3.4916 | -34.096 | -6.3602
3.4977 | -34.0961 | -6.3715
3.5038 | -34.0963 | -6.3824
3.5099 | -34.0964 | -6.3927
3.5161 | -34.0966 | -6.4031
3.5222 | -34.0967 | -6.4135
3.5283 | -34.0968 | -6.4244
3.5344 | -34.0969 | -6.4352
3.5405 | -34.0971 | -6.4458
3.5466 | -34.0973 | -6.4567
3.5527 | -34.0974 | -6.4675
3.5588 | -34.0975 | -6.4782
3.5649 | -34.0975 | -6.4888
3.571 | -34.0977 | -6.4995
3.5771 | -34.0978 | -6.5106
3.5832 | -34.098 | -6.5206
3.5893 | -34.0982 | -6.5312
3.5954 | -34.0983 | -6.5421
3.6015 | -34.0985 | -6.553
3.6076 | -34.0987 | -6.5636
3.6137 | -34.0989 | -6.5737
3.6198 | -34.0991 | -6.584
3.6259 | -34.0993 | -6.5949
3.632 | -34.0994 | -6.6056
3.6381 | -34.0996 | -6.6164
3.6442 | -34.0998 | -6.627
3.6503 | -34.1 | -6.6372
3.6565 | -34.1002 | -6.6479
3.6626 | -34.1003 | -6.6586
3.6687 | -34.1005 | -6.6693
3.6748 | -34.1008 | -6.6798
3.6809 | -34.101 | -6.6904
3.687 | -34.1012 | -6.7014
3.6931 | -34.1014 | -6.7117
3.6992 | -34.1016 | -6.7223
3.7053 | -34.1018 | -6.7328
3.7114 | -34.102 | -6.7436
3.7175 | -34.1022 | -6.7543
3.7236 | -34.1025 | -6.7644
3.7297 | -34.1027 | -6.7752
3.7358 | -34.1029 | -6.7861
3.7419 | -34.1031 | -6.797
3.748 | -34.1033 | -6.8075
3.7541 | -34.1035 | -6.8182
3.7602 | -34.1038 | -6.8286
3.7663 | -34.104 | -6.8391
3.7724 | -34.1042 | -6.8498
3.7785 | -34.1045 | -6.8602
3.7846 | -34.1047 | -6.8707
3.7907 | -34.105 | -6.8816
3.7969 | -34.1052 | -6.8924
3.803 | -34.1055 | -6.9026
3.8091 | -34.1057 | -6.9132
3.8152 | -34.106 | -6.9237
3.8213 | -34.1062 | -6.9346
3.8274 | -34.1065 | -6.9449
3.8335 | -34.1068 | -6.9556
3.8396 | -34.1071 | -6.9663
3.8457 | -34.1074 | -6.977
3.8518 | -34.1077 | -6.9876
3.8579 | -34.1081 | -6.9981
3.864 | -34.1083 | -7.0089
3.8701 | -34.1086 | -7.0199
3.8762 | -34.1089 | -7.0304
3.8823 | -34.1091 | -7.0412
3.8884 | -34.1095 | -7.0519
3.8945 | -34.1098 | -7.0628
3.9006 | -34.1101 | -7.0738
3.9067 | -34.1104 | -7.0848
3.9128 | -34.1107 | -7.0956
3.9189 | -34.111 | -7.1061
3.925 | -34.1112 | -7.1169
3.9311 | -34.1115 | -7.1271
3.9372 | -34.1118 | -7.138
3.9434 | -34.1121 | -7.1487
3.9495 | -34.1125 | -7.1595
3.9556 | -34.1128 | -7.17
3.9617 | -34.1131 | -7.1804
3.9678 | -34.1134 | -7.1908
3.9739 | -34.1137 | -7.2015
3.98 | -34.114 | -7.2123
3.9861 | -34.1143 | -7.2229
3.9922 | -34.1145 | -7.2336
3.9983 | -34.1148 | -7.2439
4.0044 | -34.1151 | -7.2542
4.0105 | -34.1154 | -7.2648
4.0166 | -34.1157 | -7.2756
4.0227 | -34.116 | -7.2862
4.0288 | -34.1163 | -7.2973
4.0349 | -34.1165 | -7.3078
4.041 | -34.1168 | -7.3184
4.0471 | -34.117 | -7.3291
4.0532 | -34.1173 | -7.3394
4.0593 | -34.1177 | -7.3503
4.0654 | -34.118 | -7.3607
4.0715 | -34.1183 | -7.3718
4.0776 | -34.1186 | -7.3824
4.0838 | -34.1189 | -7.3934
4.0899 | -34.1192 | -7.4038
4.096 | -34.1195 | -7.4144
4.1021 | -34.1197 | -7.4248
4.1082 | -34.12 | -7.4352
4.1143 | -34.1202 | -7.4455
4.1204 | -34.1206 | -7.4559
4.1265 | -34.1209 | -7.4662
4.1326 | -34.1211 | -7.4765
4.1387 | -34.1215 | -7.4869
4.1448 | -34.1217 | -7.4976
4.1509 | -34.122 | -7.5082
4.157 | -34.1222 | -7.5191
4.1631 | -34.1225 | -7.5299
4.1692 | -34.1228 | -7.5403
4.1753 | -34.1231 | -7.5508
4.1814 | -34.1233 | -7.5611
4.1875 | -34.1236 | -7.5719
4.1936 | -34.1239 | -7.5826
4.1997 | -34.1241 | -7.5928
4.2058 | -34.1244 | -7.6036
4.2119 | -34.1246 | -7.614
4.218 | -34.1248 | -7.6242
4.2241 | -34.1251 | -7.6346
4.2303 | -34.1254 | -7.6449
4.2364 | -34.1256 | -7.6554
4.2425 | -34.1258 | -7.6659
4.2486 | -34.1261 | -7.6766
4.2547 | -34.1262 | -7.6869
4.2608 | -34.1265 | -7.6975
4.2669 | -34.1267 | -7.7076
4.273 | -34.1269 | -7.7183
4.2791 | -34.1272 | -7.7288
4.2852 | -34.1274 | -7.7394
4.2913 | -34.1276 | -7.7496
4.2974 | -34.1278 | -7.7603
4.3035 | -34.1281 | -7.771
4.3096 | -34.1283 | -7.7817
4.3157 | -34.1285 | -7.792
4.3218 | -34.1288 | -7.8023
4.3279 | -34.129 | -7.8128
4.334 | -34.1292 | -7.8232
4.3401 | -34.1294 | -7.8332
4.3462 | -34.1297 | -7.8438
4.3523 | -34.13 | -7.8538
4.3584 | -34.1301 | -7.8638
4.3645 | -34.1303 | -7.8741
4.3707 | -34.1306 | -7.8843
4.3768 | -34.1308 | -7.8946
4.3829 | -34.131 | -7.9054
4.389 | -34.1312 | -7.9157
4.3951 | -34.1315 | -7.9259
4.4012 | -34.1316 | -7.9366
4.4073 | -34.1318 | -7.9468
4.4134 | -34.1321 | -7.957
4.4195 | -34.1323 | -7.9673
4.4256 | -34.1326 | -7.9775
4.4317 | -34.1328 | -7.9878
4.4378 | -34.133 | -7.9978
4.4439 | -34.1332 | -8.0079
4.45 | -34.1334 | -8.0182
4.4561 | -34.1335 | -8.0285
4.4622 | -34.1338 | -8.0389
4.4683 | -34.134 | -8.0491
4.4744 | -34.1342 | -8.0597
4.4805 | -34.1344 | -8.0696
4.4866 | -34.1346 | -8.0799
4.4927 | -34.1349 | -8.0905
4.4988 | -34.1351 | -8.1007
4.5049 | -34.1353 | -8.1106
4.511 | -34.1355 | -8.1207
4.5172 | -34.1357 | -8.1309
4.5233 | -34.1359 | -8.1415
4.5294 | -34.1362 | -8.1517
4.5355 | -34.1364 | -8.162
4.5416 | -34.1366 | -8.1722
4.5477 | -34.1368 | -8.1824
4.5538 | -34.137 | -8.1923
4.5599 | -34.1373 | -8.2025
4.566 | -34.1375 | -8.2127
4.5721 | -34.1377 | -8.223
4.5782 | -34.1379 | -8.2335
4.5843 | -34.1381 | -8.2438
4.5904 | -34.1384 | -8.2541
4.5965 | -34.1387 | -8.2642
4.6026 | -34.1389 | -8.2749
4.6087 | -34.1391 | -8.2851
4.6148 | -34.1393 | -8.2956
4.6209 | -34.1395 | -8.3058
4.627 | -34.1398 | -8.3163
4.6331 | -34.1399 | -8.3264
4.6392 | -34.1401 | -8.3368
4.6453 | -34.1404 | -8.347
4.6514 | -34.1406 | -8.3573
4.6576 | -34.1408 | -8.3679
4.6637 | -34.141 | -8.3779
4.6698 | -34.1412 | -8.388
4.6759 | -34.1414 | -8.3986
4.682 | -34.1416 | -8.4087
4.6881 | -34.1417 | -8.419
4.6942 | -34.142 | -8.4292
4.7003 | -34.1422 | -8.4396
4.7064 | -34.1424 | -8.4496
4.7125 | -34.1427 | -8.4601
4.7186 | -34.1429 | -8.4702
4.7247 | -34.1431 | -8.4805
4.7308 | -34.1434 | -8.4907
4.7369 | -34.1436 | -8.5011
4.743 | -34.1438 | -8.5117
4.7491 | -34.144 | -8.5223
4.7552 | -34.1442 | -8.5327
4.7613 | -34.1444 | -8.543
4.7674 | -34.1446 | -8.5532
4.7735 | -34.1448 | -8.5639
4.7796 | -34.1451 | -8.5745
4.7857 | -34.1453 | -8.5849
4.7918 | -34.1456 | -8.595
4.7979 | -34.1458 | -8.6052
4.8041 | -34.146 | -8.6156
4.8102 | -34.1463 | -8.6256
4.8163 | -34.1465 | -8.6358
4.8224 | -34.1467 | -8.6462
4.8285 | -34.147 | -8.6562
4.8346 | -34.1472 | -8.6663
4.8407 | -34.1474 | -8.6767
4.8468 | -34.1477 | -8.6873
4.8529 | -34.1479 | -8.6979
4.859 | -34.1481 | -8.7085
4.8651 | -34.1484 | -8.7186
4.8712 | -34.1486 | -8.7287
4.8773 | -34.1488 | -8.7391
4.8834 | -34.149 | -8.7494
4.8895 | -34.1493 | -8.7595
4.8956 | -34.1495 | -8.7699
4.9017 | -34.1497 | -8.78
4.9078 | -34.1499 | -8.7906
4.9139 | -34.1501 | -8.8005
4.92 | -34.1504 | -8.8106
4.9261 | -34.1507 | -8.8205
4.9322 | -34.1509 | -8.8309
4.9383 | -34.1512 | -8.8412
4.9445 | -34.1514 | -8.8515
4.9506 | -34.1517 | -8.8616
4.9567 | -34.1519 | -8.8717
4.9628 | -34.1521 | -8.882
4.9689 | -34.1524 | -8.8926
4.975 | -34.1526 | -8.9032
4.9811 | -34.1528 | -8.9132
4.9872 | -34.1531 | -8.9234
4.9933 | -34.1534 | -8.934
4.9994 | -34.1537 | -8.9446
5.0055 | -34.1539 | -8.955
5.0116 | -34.1542 | -8.9657
5.0177 | -34.1545 | -8.9763
5.0238 | -34.1548 | -8.9869
5.0299 | -34.1551 | -8.9976
5.036 | -34.1554 | -9.0079
5.0421 | -34.1556 | -9.0182
5.0482 | -34.1559 | -9.0285
5.0543 | -34.1561 | -9.0386
5.0604 | -34.1564 | -9.0489
5.0665 | -34.1567 | -9.059
5.0726 | -34.1569 | -9.0697
5.0787 | -34.1572 | -9.0803
5.0848 | -34.1575 | -9.0907
5.091 | -34.1577 | -9.1008
5.0971 | -34.158 | -9.1109
5.1032 | -34.1583 | -9.1216
5.1093 | -34.1586 | -9.1319
5.1154 | -34.1588 | -9.142
5.1215 | -34.1591 | -9.1522
5.1276 | -34.1593 | -9.1623
5.1337 | -34.1597 | -9.1724
5.1398 | -34.1599 | -9.1825
5.1459 | -34.1602 | -9.1931
5.152 | -34.1605 | -9.2036
5.1581 | -34.1607 | -9.2141
5.1642 | -34.161 | -9.2246
5.1703 | -34.1613 | -9.2349
5.1764 | -34.1616 | -9.2451
5.1825 | -34.1619 | -9.2552
5.1886 | -34.1621 | -9.2653
5.1947 | -34.1624 | -9.2757
5.2008 | -34.1627 | -9.2862
5.2069 | -34.163 | -9.2968
5.213 | -34.1633 | -9.3069
5.2191 | -34.1635 | -9.3173
5.2252 | -34.1638 | -9.3277
5.2314 | -34.1641 | -9.3383
5.2375 | -34.1643 | -9.3484
5.2436 | -34.1646 | -9.3588
5.2497 | -34.1649 | -9.369
5.2558 | -34.1652 | -9.3796
5.2619 | -34.1655 | -9.3903
5.268 | -34.1658 | -9.4008
5.2741 | -34.1661 | -9.4112
5.2802 | -34.1664 | -9.422
5.2863 | -34.1667 | -9.4323
5.2924 | -34.167 | -9.4428
5.2985 | -34.1673 | -9.4533
5.3046 | -34.1676 | -9.4637
5.3107 | -34.1679 | -9.4744
5.3168 | -34.1682 | -9.4849
5.3229 | -34.1684 | -9.4953
5.329 | -34.1687 | -9.5057
5.3351 | -34.169 | -9.5158
5.3412 | -34.1693 | -9.5262
5.3473 | -34.1696 | -9.5367
5.3534 | -34.1699 | -9.5472
5.3595 | -34.1702 | -9.5577
5.3656 | -34.1705 | -9.5681
5.3717 | -34.1708 | -9.5789
5.3779 | -34.1711 | -9.5897
5.384 | -34.1714 | -9.6001
5.3901 | -34.1717 | -9.6108
5.3962 | -34.172 | -9.621
5.4023 | -34.1723 | -9.6317
5.4084 | -34.1726 | -9.6421
5.4145 | -34.173 | -9.6523
5.4206 | -34.1732 | -9.6628
5.4267 | -34.1735 | -9.6735
5.4328 | -34.1738 | -9.6841
5.4389 | -34.1741 | -9.6943
5.445 | -34.1744 | -9.7051
5.4511 | -34.1747 | -9.7153
5.4572 | -34.175 | -9.7258
5.4633 | -34.1753 | -9.7363
5.4694 | -34.1756 | -9.7468
5.4755 | -34.1759 | -9.757
5.4816 | -34.1762 | -9.7675
5.4877 | -34.1765 | -9.7778
5.4938 | -34.1768 | -9.788
5.4999 | -34.1771 | -9.7984
5.506 | -34.1775 | -9.8089
5.5121 | -34.1778 | -9.8192
5.5183 | -34.1781 | -9.8295
5.5244 | -34.1783 | -9.8402
5.5305 | -34.1787 | -9.8506
5.5366 | -34.1789 | -9.8608
5.5427 | -34.1792 | -9.8714
5.5488 | -34.1795 | -9.8819
5.5549 | -34.1798 | -9.8924
5.561 | -34.1801 | -9.9031
5.5671 | -34.1804 | -9.9134
5.5732 | -34.1807 | -9.9241
5.5793 | -34.1811 | -9.9349
5.5854 | -34.1814 | -9.9453
5.5915 | -34.1817 | -9.9561
5.5976 | -34.182 | -9.9667
5.6037 | -34.1823 | -9.9772
5.6098 | -34.1825 | -9.9876
5.6159 | -34.1829 | -9.998
5.622 | -34.1832 | -10.0085
5.6281 | -34.1835 | -10.0191
5.6342 | -34.1838 | -10.0299
5.6403 | -34.1841 | -10.0404
5.6464 | -34.1844 | -10.051
5.6525 | -34.1847 | -10.0616
5.6586 | -34.185 | -10.0718
5.6648 | -34.1854 | -10.0823
5.6709 | -34.1857 | -10.093
5.677 | -34.186 | -10.1035
5.6831 | -34.1863 | -10.114
5.6892 | -34.1866 | -10.1247
5.6953 | -34.1869 | -10.1352
5.7014 | -34.1872 | -10.1459
5.7075 | -34.1875 | -10.1563
5.7136 | -34.1878 | -10.1668
5.7197 | -34.1881 | -10.1773
5.7258 | -34.1884 | -10.1876
5.7319 | -34.1887 | -10.1981
5.738 | -34.189 | -10.2086
5.7441 | -34.1894 | -10.2191
5.7502 | -34.1897 | -10.2293
5.7563 | -34.19 | -10.2398
5.7624 | -34.1903 | -10.2505
5.7685 | -34.1907 | -10.261
5.7746 | -34.1909 | -10.2717
5.7807 | -34.1912 | -10.2823
5.7868 | -34.1915 | -10.293
5.7929 | -34.1918 | -10.3034
5.799 | -34.1921 | -10.3138
5.8052 | -34.1924 | -10.3241
5.8113 | -34.1927 | -10.3345
5.8174 | -34.193 | -10.3449
5.8235 | -34.1933 | -10.3553
5.8296 | -34.1936 | -10.366
5.8357 | -34.1939 | -10.3765
5.8418 | -34.1942 | -10.3871
5.8479 | -34.1944 | -10.3977
5.854 | -34.1948 | -10.4083
5.8601 | -34.1951 | -10.4186
5.8662 | -34.1954 | -10.4293
5.8723 | -34.1957 | -10.4396
5.8784 | -34.1959 | -10.45
5.8845 | -34.1963 | -10.4603
5.8906 | -34.1965 | -10.4709
5.8967 | -34.1968 | -10.4812
5.9028 | -34.1972 | -10.4915
5.9089 | -34.1975 | -10.5018
5.915 | -34.1978 | -10.5121
5.9211 | -34.1981 | -10.5225
5.9272 | -34.1984 | -10.5327
5.9333 | -34.1987 | -10.5431
5.9394 | -34.199 | -10.5535
5.9455 | -34.1994 | -10.5637
5.9517 | -34.1997 | -10.574
5.9578 | -34.2001 | -10.5846
5.9639 | -34.2004 | -10.5946
5.97 | -34.2007 | -10.6051
5.9761 | -34.201 | -10.6155
5.9822 | -34.2013 | -10.626
5.9883 | -34.2016 | -10.6364
5.9944 | -34.202 | -10.6469
6.0005 | -34.2023 | -10.6575
6.0066 | -34.2026 | -10.6677
6.0127 | -34.2029 | -10.6778
6.0188 | -34.2032 | -10.6885
6.0249 | -34.2035 | -10.699
6.031 | -34.2038 | -10.7093
6.0371 | -34.2042 | -10.7197
6.0432 | -34.2045 | -10.7299
6.0493 | -34.2048 | -10.7401
6.0554 | -34.2051 | -10.7504
6.0615 | -34.2054 | -10.7607
6.0676 | -34.2058 | -10.7713
6.0737 | -34.2061 | -10.7818
6.0798 | -34.2064 | -10.7922
6.0859 | -34.2067 | -10.8025
6.0921 | -34.207 | -10.8126
6.0982 | -34.2074 | -10.823
6.1043 | -34.2077 | -10.8332
6.1104 | -34.208 | -10.8438
6.1165 | -34.2084 | -10.8537
6.1226 | -34.2087 | -10.864
6.1287 | -34.209 | -10.8743
6.1348 | -34.2093 | -10.8846
6.1409 | -34.2097 | -10.8949
6.147 | -34.21 | -10.9052
6.1531 | -34.2103 | -10.9154
6.1592 | -34.2106 | -10.9257
6.1653 | -34.2109 | -10.9359
6.1714 | -34.2112 | -10.9462
6.1775 | -34.2115 | -10.9566
6.1836 | -34.2118 | -10.967
6.1897 | -34.2122 | -10.9772
6.1958 | -34.2125 | -10.9878
6.2019 | -34.2128 | -10.9982
6.208 | -34.2132 | -11.0085
6.2141 | -34.2135 | -11.0188
6.2202 | -34.2138 | -11.0292
6.2263 | -34.2142 | -11.0394
6.2325 | -34.2144 | -11.0494
6.2386 | -34.2148 | -11.0594
6.2447 | -34.2151 | -11.0699
6.2508 | -34.2154 | -11.0802
6.2569 | -34.2158 | -11.0904
6.263 | -34.2161 | -11.1007
6.2691 | -34.2164 | -11.1109
6.2752 | -34.2167 | -11.1212
6.2813 | -34.217 | -11.1314
6.2874 | -34.2173 | -11.1416
6.2935 | -34.2177 | -11.1518
6.2996 | -34.2179 | -11.162
6.3057 | -34.2183 | -11.1723
6.3118 | -34.2185 | -11.1827
6.3179 | -34.2189 | -11.1929
6.324 | -34.2192 | -11.2031
6.3301 | -34.2195 | -11.2133
6.3362 | -34.2199 | -11.2235
6.3423 | -34.2202 | -11.2338
6.3484 | -34.2205 | -11.2442
6.3545 | -34.2208 | -11.2541
6.3606 | -34.2211 | -11.2641
6.3667 | -34.2214 | -11.2741
6.3728 | -34.2217 | -11.2844
6.379 | -34.2221 | -11.2946
6.3851 | -34.2224 | -11.3046
6.3912 | -34.2227 | -11.3148
6.3973 | -34.223 | -11.325
6.4034 | -34.2233 | -11.335
6.4095 | -34.2236 | -11.3452
6.4156 | -34.2239 | -11.3554
6.4217 | -34.2242 | -11.3655
6.4278 | -34.2245 | -11.3758
6.4339 | -34.2249 | -11.3859
6.44 | -34.2252 | -11.396
6.4461 | -34.2255 | -11.4064
6.4522 | -34.2258 | -11.4164
6.4583 | -34.2261 | -11.4267
6.4644 | -34.2264 | -11.437
6.4705 | -34.2267 | -11.4468
6.4766 | -34.2271 | -11.457
6.4827 | -34.2274 | -11.4672
6.4888 | -34.2277 | -11.4772
6.4949 | -34.228 | -11.4871
6.501 | -34.2283 | -11.4973
6.5071 | -34.2286 | -11.5072
6.5132 | -34.2289 | -11.5174
6.5194 | -34.2292 | -11.5276
6.5255 | -34.2296 | -11.5375
6.5316 | -34.2298 | -11.5478
6.5377 | -34.2302 | -11.5579
6.5438 | -34.2304 | -11.5678
6.5499 | -34.2308 | -11.5778
6.556 | -34.2311 | -11.5879
6.5621 | -34.2314 | -11.5983
6.5682 | -34.2317 | -11.6084
6.5743 | -34.2321 | -11.6184
6.5804 | -34.2324 | -11.6287
6.5865 | -34.2327 | -11.6388
6.5926 | -34.233 | -11.6488
6.5987 | -34.2333 | -11.6586
6.6048 | -34.2336 | -11.6688
6.6109 | -34.234 | -11.6788
6.617 | -34.2342 | -11.689
6.6231 | -34.2345 | -11.6991
6.6292 | -34.2349 | -11.7092
6.6353 | -34.2352 | -11.7195
6.6414 | -34.2355 | -11.7295
6.6475 | -34.2359 | -11.7396
6.6536 | -34.2361 | -11.7495
6.6597 | -34.2365 | -11.7593
6.6659 | -34.2368 | -11.7694
6.672 | -34.2371 | -11.7794
6.6781 | -34.2374 | -11.7895
6.6842 | -34.2377 | -11.7994
6.6903 | -34.2381 | -11.8094
6.6964 | -34.2384 | -11.8195
6.7025 | -34.2387 | -11.8294
6.7086 | -34.239 | -11.8394
6.7147 | -34.2393 | -11.8495
6.7208 | -34.2396 | -11.8597
6.7269 | -34.24 | -11.8697
6.733 | -34.2402 | -11.8795
6.7391 | -34.2406 | -11.8895
6.7452 | -34.2409 | -11.8994
6.7513 | -34.2412 | -11.9091
6.7574 | -34.2415 | -11.9192
6.7635 | -34.2419 | -11.929
6.7696 | -34.2422 | -11.9389
6.7757 | -34.2425 | -11.949
6.7818 | -34.2429 | -11.959
6.7879 | -34.2432 | -11.9688
6.794 | -34.2435 | -11.9786
6.8001 | -34.2438 | -11.9885
6.8063 | -34.2441 | -11.9985
6.8124 | -34.2444 | -12.0086
6.8185 | -34.2447 | -12.0188
6.8246 | -34.245 | -12.0285
6.8307 | -34.2453 | -12.0384
6.8368 | -34.2457 | -12.0486
6.8429 | -34.246 | -12.0586
6.849 | -34.2463 | -12.0686
6.8551 | -34.2467 | -12.0785
6.8612 | -34.247 | -12.0884
6.8673 | -34.2473 | -12.0983
6.8734 | -34.2476 | -12.1084
6.8795 | -34.2479 | -12.1184
6.8856 | -34.2483 | -12.1283
6.8917 | -34.2486 | -12.1382
6.8978 | -34.2489 | -12.1479
6.9039 | -34.2492 | -12.1576
6.91 | -34.2495 | -12.1676
6.9161 | -34.2499 | -12.1773
6.9222 | -34.2502 | -12.1873
6.9283 | -34.2505 | -12.1969
6.9344 | -34.2509 | -12.2069
6.9405 | -34.2512 | -12.2168
6.9466 | -34.2516 | -12.2269
6.9528 | -34.2519 | -12.237
6.9589 | -34.2522 | -12.2471
6.965 | -34.2526 | -12.2571
6.9711 | -34.2529 | -12.2671
6.9772 | -34.2533 | -12.277
6.9833 | -34.2536 | -12.2869
6.9894 | -34.2539 | -12.297
6.9955 | -34.2542 | -12.3067
7.0016 | -34.2546 | -12.3166
7.0077 | -34.2549 | -12.3265
7.0138 | -34.2552 | -12.3363
7.0199 | -34.2556 | -12.3461
7.026 | -34.2559 | -12.356
7.0321 | -34.2562 | -12.3657
7.0382 | -34.2566 | -12.3753
7.0443 | -34.2569 | -12.3854
7.0504 | -34.2572 | -12.3954
7.0565 | -34.2576 | -12.4053
7.0626 | -34.2579 | -12.4149
7.0687 | -34.2582 | -12.4244
7.0748 | -34.2586 | -12.4341
7.0809 | -34.2589 | -12.4439
7.087 | -34.2593 | -12.4538
7.0932 | -34.2596 | -12.4637
7.0993 | -34.2599 | -12.4735
7.1054 | -34.2603 | -12.4833
7.1115 | -34.2606 | -12.4929
7.1176 | -34.2609 | -12.5027
7.1237 | -34.2612 | -12.5125
7.1298 | -34.2616 | -12.5222
7.1359 | -34.2619 | -12.5321
7.142 | -34.2622 | -12.5417
7.1481 | -34.2626 | -12.5514
7.1542 | -34.2629 | -12.5613
7.1603 | -34.2632 | -12.5711
7.1664 | -34.2636 | -12.5809
7.1725 | -34.2639 | -12.5908
7.1786 | -34.2643 | -12.6006
7.1847 | -34.2646 | -12.6105
7.1908 | -34.2649 | -12.6204
7.1969 | -34.2652 | -12.6304
7.203 | -34.2656 | -12.6401
7.2091 | -34.266 | -12.6498
7.2152 | -34.2663 | -12.6597
7.2213 | -34.2666 | -12.6694
7.2274 | -34.267 | -12.6791
7.2335 | -34.2673 | -12.6889
7.2397 | -34.2677 | -12.6986
7.2458 | -34.268 | -12.7082
7.2519 | -34.2684 | -12.718
7.258 | -34.2687 | -12.728
7.2641 | -34.269 | -12.7378
7.2702 | -34.2694 | -12.7478
7.2763 | -34.2697 | -12.7576
7.2824 | -34.2701 | -12.7675
7.2885 | -34.2704 | -12.7774
7.2946 | -34.2708 | -12.7872
7.3007 | -34.2711 | -12.7971
7.3068 | -34.2714 | -12.8071
7.3129 | -34.2717 | -12.8171
7.319 | -34.2721 | -12.8269
7.3251 | -34.2724 | -12.8366
7.3312 | -34.2728 | -12.8464
7.3373 | -34.2731 | -12.856
7.3434 | -34.2735 | -12.8657
7.3495 | -34.2739 | -12.8754
7.3556 | -34.2742 | -12.8853
7.3617 | -34.2746 | -12.8952
7.3678 | -34.2749 | -12.905
7.3739 | -34.2752 | -12.915
7.3801 | -34.2756 | -12.925
7.3862 | -34.2759 | -12.9348
7.3923 | -34.2763 | -12.9447
7.3984 | -34.2766 | -12.9543
7.4045 | -34.277 | -12.964
7.4106 | -34.2774 | -12.9738
7.4167 | -34.2777 | -12.9837
7.4228 | -34.2781 | -12.9936
7.4289 | -34.2784 | -13.0034
7.435 | -34.2788 | -13.013
7.4411 | -34.2791 | -13.0229
7.4472 | -34.2794 | -13.0327
7.4533 | -34.2798 | -13.0424
7.4594 | -34.2802 | -13.0524
7.4655 | -34.2805 | -13.0621
7.4716 | -34.2809 | -13.072
7.4777 | -34.2812 | -13.0818
7.4838 | -34.2816 | -13.0915
7.4899 | -34.2819 | -13.1013
7.496 | -34.2822 | -13.1109
7.5021 | -34.2826 | -13.1207
7.5082 | -34.2829 | -13.1305
7.5143 | -34.2833 | -13.1404
7.5204 | -34.2836 | -13.1503
7.5266 | -34.284 | -13.16
7.5327 | -34.2843 | -13.1697
7.5388 | -34.2847 | -13.1798
7.5449 | -34.285 | -13.1898
7.551 | -34.2854 | -13.1995
7.5571 | -34.2858 | -13.2094
7.5632 | -34.2861 | -13.2193
7.5693 | -34.2865 | -13.229
7.5754 | -34.2869 | -13.2387
7.5815 | -34.2872 | -13.2485
7.5876 | -34.2875 | -13.2586
7.5937 | -34.2879 | -13.2685
7.5998 | -34.2883 | -13.2783
7.6059 | -34.2887 | -13.2883
7.612 | -34.289 | -13.2981
7.6181 | -34.2894 | -13.3078
7.6242 | -34.2897 | -13.3179
7.6303 | -34.2901 | -13.328
7.6364 | -34.2904 | -13.3379
7.6425 | -34.2908 | -13.3477
7.6486 | -34.2911 | -13.3576
7.6547 | -34.2915 | -13.3673
7.6608 | -34.2918 | -13.3771
7.667 | -34.2922 | -13.3871
7.6731 | -34.2925 | -13.397
7.6792 | -34.2929 | -13.407
7.6853 | -34.2933 | -13.417
7.6914 | -34.2936 | -13.427
7.6975 | -34.294 | -13.4369
7.7036 | -34.2944 | -13.4468
7.7097 | -34.2947 | -13.4568
7.7158 | -34.2951 | -13.4668
7.7219 | -34.2955 | -13.4765
7.728 | -34.2958 | -13.4865
7.7341 | -34.2962 | -13.4965
7.7402 | -34.2965 | -13.5064
7.7463 | -34.2969 | -13.5163
7.7524 | -34.2972 | -13.5265
7.7585 | -34.2976 | -13.5364
7.7646 | -34.298 | -13.5463
7.7707 | -34.2983 | -13.5564
7.7768 | -34.2986 | -13.5664
7.7829 | -34.299 | -13.5766
7.789 | -34.2994 | -13.5867
7.7951 | -34.2997 | -13.5966
7.8012 | -34.3001 | -13.6065
7.8073 | -34.3005 | -13.6166
7.8135 | -34.3008 | -13.6265
7.8196 | -34.3012 | -13.6367
7.8257 | -34.3016 | -13.6468
7.8318 | -34.3019 | -13.6568
7.8379 | -34.3023 | -13.6666
7.844 | -34.3026 | -13.6765
7.8501 | -34.303 | -13.6866
7.8562 | -34.3034 | -13.6967
7.8623 | -34.3038 | -13.7066
7.8684 | -34.3041 | -13.7166
7.8745 | -34.3045 | -13.7266
7.8806 | -34.3048 | -13.7367
7.8867 | -34.3052 | -13.7464
7.8928 | -34.3056 | -13.7562
7.8989 | -34.3059 | -13.7661
7.905 | -34.3063 | -13.7762
7.9111 | -34.3066 | -13.7862
7.9172 | -34.307 | -13.7961
7.9233 | -34.3073 | -13.8062
7.9294 | -34.3077 | -13.8162
7.9355 | -34.3081 | -13.8259
7.9416 | -34.3084 | -13.8359
7.9477 | -34.3088 | -13.8458
7.9539 | -34.3091 | -13.8557
7.96 | -34.3095 | -13.8658
7.9661 | -34.3099 | -13.8758
7.9722 | -34.3103 | -13.8856
7.9783 | -34.3106 | -13.8958
7.9844 | -34.311 | -13.9058
7.9905 | -34.3114 | -13.9159
7.9966 | -34.3117 | -13.926
8.0027 | -34.3121 | -13.936
8.0088 | -34.3124 | -13.9459
8.0149 | -34.3128 | -13.9559
8.021 | -34.3131 | -13.966
8.0271 | -34.3135 | -13.9762
8.0332 | -34.3139 | -13.9862
8.0393 | -34.3143 | -13.9963
8.0454 | -34.3146 | -14.0065
8.0515 | -34.315 | -14.0165
8.0576 | -34.3154 | -14.0267
8.0637 | -34.3157 | -14.0367
8.0698 | -34.3161 | -14.0465
8.0759 | -34.3164 | -14.0564
8.082 | -34.3168 | -14.0663
8.0881 | -34.3171 | -14.0763
8.0942 | -34.3175 | -14.0864
8.1004 | -34.3179 | -14.0965
8.1065 | -34.3183 | -14.1063
8.1126 | -34.3186 | -14.1163
8.1187 | -34.319 | -14.1264
8.1248 | -34.3193 | -14.1362
8.1309 | -34.3197 | -14.1461
8.137 | -34.32 | -14.1559
8.1431 | -34.3204 | -14.1661
8.1492 | -34.3208 | -14.176
8.1553 | -34.3211 | -14.1861
8.1614 | -34.3215 | -14.1962
8.1675 | -34.3218 | -14.2062
8.1736 | -34.3222 | -14.2164
8.1797 | -34.3225 | -14.2264
8.1858 | -34.3229 | -14.2365
8.1919 | -34.3233 | -14.2465
8.198 | -34.3236 | -14.2563
8.2041 | -34.324 | -14.2663
8.2102 | -34.3244 | -14.2763
8.2163 | -34.3248 | -14.2861
8.2224 | -34.3251 | -14.2961
8.2285 | -34.3255 | -14.3058
8.2346 | -34.3258 | -14.3158
8.2408 | -34.3262 | -14.3256
8.2469 | -34.3266 | -14.3357
8.253 | -34.327 | -14.3455
8.2591 | -34.3273 | -14.3555
8.2652 | -34.3278 | -14.3655
8.2713 | -34.3281 | -14.3756
8.2774 | -34.3285 | -14.3857
8.2835 | -34.3289 | -14.3959
8.2896 | -34.3292 | -14.4057
8.2957 | -34.3296 | -14.4157
8.3018 | -34.33 | -14.4257
8.3079 | -34.3304 | -14.4355
8.314 | -34.3307 | -14.4455
8.3201 | -34.3311 | -14.4554
8.3262 | -34.3315 | -14.4654
8.3323 | -34.3318 | -14.4754
8.3384 | -34.3322 | -14.4855
8.3445 | -34.3326 | -14.4954
8.3506 | -34.3329 | -14.5054
8.3567 | -34.3333 | -14.5155
8.3628 | -34.3337 | -14.5254
8.3689 | -34.334 | -14.5353
8.375 | -34.3344 | -14.5454
8.3812 | -34.3348 | -14.5552
8.3873 | -34.3352 | -14.5652
8.3934 | -34.3355 | -14.5753
8.3995 | -34.3359 | -14.5851
8.4056 | -34.3362 | -14.595
8.4117 | -34.3366 | -14.6051
8.4178 | -34.337 | -14.6151
8.4239 | -34.3373 | -14.6249
8.43 | -34.3377 | -14.6349
8.4361 | -34.338 | -14.6449
8.4422 | -34.3384 | -14.6547
8.4483 | -34.3388 | -14.6646
8.4544 | -34.3392 | -14.6745
8.4605 | -34.3395 | -14.6844
8.4666 | -34.3399 | -14.6945
8.4727 | -34.3402 | -14.7045
8.4788 | -34.3406 | -14.7144
8.4849 | -34.341 | -14.7244
8.491 | -34.3413 | -14.7343
8.4971 | -34.3417 | -14.744
8.5032 | -34.3421 | -14.7538
8.5093 | -34.3424 | -14.7636
8.5154 | -34.3428 | -14.7736
8.5215 | -34.3432 | -14.7834
8.5277 | -34.3435 | -14.7934
8.5338 | -34.3439 | -14.8032
8.5399 | -34.3443 | -14.8131
8.546 | -34.3447 | -14.8229
8.5521 | -34.345 | -14.8328
8.5582 | -34.3454 | -14.8426
8.5643 | -34.3458 | -14.8523
8.5704 | -34.3461 | -14.8619
8.5765 | -34.3465 | -14.872
8.5826 | -34.3469 | -14.8816
8.5887 | -34.3473 | -14.8914
8.5948 | -34.3476 | -14.9013
8.6009 | -34.348 | -14.911
8.607 | -34.3484 | -14.921
8.6131 | -34.3487 | -14.9309
8.6192 | -34.3491 | -14.9406
8.6253 | -34.3495 | -14.9504
8.6314 | -34.3499 | -14.9602
8.6375 | -34.3502 | -14.9701
8.6436 | -34.3506 | -14.98
8.6497 | -34.351 | -14.9897
8.6558 | -34.3513 | -14.9995
8.6619 | -34.3517 | -15.0094
8.6681 | -34.352 | -15.0193
8.6742 | -34.3524 | -15.0291
8.6803 | -34.3528 | -15.0389
8.6864 | -34.3532 | -15.0485
8.6925 | -34.3535 | -15.0581
8.6986 | -34.3539 | -15.0681
8.7047 | -34.3542 | -15.0781
8.7108 | -34.3546 | -15.0879
8.7169 | -34.355 | -15.0977
8.723 | -34.3553 | -15.1073
8.7291 | -34.3557 | -15.117
8.7352 | -34.3561 | -15.127
8.7413 | -34.3565 | -15.1369
8.7474 | -34.3568 | -15.1465
8.7535 | -34.3572 | -15.1563
8.7596 | -34.3576 | -15.1661
8.7657 | -34.358 | -15.176
8.7718 | -34.3583 | -15.1859
8.7779 | -34.3587 | -15.1958
8.784 | -34.3591 | -15.2056
8.7901 | -34.3595 | -15.2154
8.7962 | -34.3598 | -15.2249
8.8023 | -34.3602 | -15.2347
8.8084 | -34.3606 | -15.2444
8.8146 | -34.361 | -15.2542
8.8207 | -34.3613 | -15.2639
8.8268 | -34.3617 | -15.2733
8.8329 | -34.3621 | -15.283
8.839 | -34.3625 | -15.2929
8.8451 | -34.3628 | -15.3025
8.8512 | -34.3632 | -15.3121
8.8573 | -34.3636 | -15.3219
8.8634 | -34.3639 | -15.3314
8.8695 | -34.3643 | -15.3413
8.8756 | -34.3647 | -15.3511
8.8817 | -34.365 | -15.3607
8.8878 | -34.3654 | -15.3707
8.8939 | -34.3658 | -15.3804
8.9 | -34.3662 | -15.3901
8.9061 | -34.3665 | -15.3999
8.9122 | -34.3669 | -15.4096
8.9183 | -34.3673 | -15.4194
8.9244 | -34.3677 | -15.4292
8.9305 | -34.3681 | -15.4387
8.9366 | -34.3685 | -15.4485
8.9427 | -34.3688 | -15.4582
8.9488 | -34.3692 | -15.4679
8.955 | -34.3696 | -15.4776
8.9611 | -34.37 | -15.4874
8.9672 | -34.3704 | -15.4971
8.9733 | -34.3708 | -15.5066
8.9794 | -34.3712 | -15.5161
8.9855 | -34.3716 | -15.5259
8.9916 | -34.3719 | -15.5358
8.9977 | -34.3723 | -15.5455
9.0038 | -34.3728 | -15.5551
9.0099 | -34.3731 | -15.5648
9.016 | -34.3735 | -15.5746
9.0221 | -34.3739 | -15.5842
9.0282 | -34.3743 | -15.5937
9.0343 | -34.3747 | -15.6034
9.0404 | -34.3751 | -15.6131
9.0465 | -34.3755 | -15.623
9.0526 | -34.3759 | -15.6325
9.0587 | -34.3763 | -15.6422
9.0648 | -34.3767 | -15.6518
9.0709 | -34.3771 | -15.6615
9.077 | -34.3775 | -15.6712
9.0831 | -34.3779 | -15.681
9.0892 | -34.3783 | -15.6908
9.0953 | -34.3787 | -15.7006
9.1015 | -34.3791 | -15.7102
9.1076 | -34.3795 | -15.7198
9.1137 | -34.3799 | -15.7294
9.1198 | -34.3803 | -15.7391
9.1259 | -34.3807 | -15.7488
9.132 | -34.381 | -15.7586
9.1381 | -34.3815 | -15.7681
9.1442 | -34.3819 | -15.7774
9.1503 | -34.3823 | -15.7872
9.1564 | -34.3827 | -15.7968
9.1625 | -34.383 | -15.8065
9.1686 | -34.3835 | -15.8161
9.1747 | -34.3839 | -15.8257
9.1808 | -34.3843 | -15.8353
9.1869 | -34.3847 | -15.8449
9.193 | -34.3851 | -15.8545
9.1991 | -34.3855 | -15.8643
9.2052 | -34.3859 | -15.8739
9.2113 | -34.3863 | -15.8835
9.2174 | -34.3867 | -15.8931
9.2235 | -34.3871 | -15.9028
9.2296 | -34.3876 | -15.9123
9.2357 | -34.3879 | -15.9221
9.2419 | -34.3884 | -15.9317
9.248 | -34.3888 | -15.9413
9.2541 | -34.3892 | -15.951
9.2602 | -34.3896 | -15.9608
9.2663 | -34.39 | -15.9701
9.2724 | -34.3904 | -15.9797
9.2785 | -34.3908 | -15.9893
9.2846 | -34.3912 | -15.9987
9.2907 | -34.3917 | -16.0083
9.2968 | -34.3921 | -16.018
9.3029 | -34.3925 | -16.0278
9.309 | -34.3929 | -16.0374
9.3151 | -34.3933 | -16.047
9.3212 | -34.3937 | -16.0564
9.3273 | -34.3941 | -16.066
9.3334 | -34.3945 | -16.0755
9.3395 | -34.3949 | -16.0849
9.3456 | -34.3953 | -16.0944
9.3517 | -34.3957 | -16.104
9.3578 | -34.3961 | -16.1135
9.3639 | -34.3965 | -16.123
9.37 | -34.397 | -16.1326
9.3761 | -34.3974 | -16.1424
9.3822 | -34.3978 | -16.1519
9.3884 | -34.3982 | -16.1613
9.3945 | -34.3986 | -16.171
9.4006 | -34.3991 | -16.1806
9.4067 | -34.3995 | -16.1903
9.4128 | -34.3999 | -16.2
9.4189 | -34.4003 | -16.2098
9.425 | -34.4007 | -16.2192
9.4311 | -34.4011 | -16.2288
9.4372 | -34.4016 | -16.2384
9.4433 | -34.402 | -16.248
9.4494 | -34.4024 | -16.2573
9.4555 | -34.4028 | -16.267
9.4616 | -34.4033 | -16.2765
9.4677 | -34.4037 | -16.2861
9.4738 | -34.4041 | -16.2957
9.4799 | -34.4045 | -16.3052
9.486 | -34.4049 | -16.3148
9.4921 | -34.4053 | -16.3244
9.4982 | -34.4058 | -16.3337
9.5043 | -34.4062 | -16.3433
9.5104 | -34.4066 | -16.3526
9.5165 | -34.407 | -16.3622
9.5226 | -34.4074 | -16.3719
9.5288 | -34.4079 | -16.3812
9.5349 | -34.4083 | -16.3908
9.541 | -34.4087 | -16.4002
9.5471 | -34.4091 | -16.4097
9.5532 | -34.4095 | -16.4193
9.5593 | -34.41 | -16.4289
9.5654 | -34.4104 | -16.4383
9.5715 | -34.4108 | -16.4478
9.5776 | -34.4112 | -16.4575
9.5837 | -34.4116 | -16.4671
9.5898 | -34.412 | -16.4767
9.5959 | -34.4124 | -16.4862
9.602 | -34.4128 | -16.4956
9.6081 | -34.4133 | -16.5051
9.6142 | -34.4137 | -16.5145
9.6203 | -34.4141 | -16.524
9.6264 | -34.4145 | -16.5335
9.6325 | -34.4149 | -16.5431
9.6386 | -34.4153 | -16.5525
9.6447 | -34.4157 | -16.5621
9.6508 | -34.4162 | -16.5715
9.6569 | -34.4166 | -16.581
9.663 | -34.417 | -16.5906
9.6691 | -34.4174 | -16.6001
9.6753 | -34.4179 | -16.6096
9.6814 | -34.4183 | -16.619
9.6875 | -34.4187 | -16.6284
9.6936 | -34.4191 | -16.6378
9.6997 | -34.4195 | -16.6474
9.7058 | -34.42 | -16.6568
9.7119 | -34.4204 | -16.6664
9.718 | -34.4208 | -16.6761
9.7241 | -34.4213 | -16.6855
9.7302 | -34.4217 | -16.695
9.7363 | -34.4221 | -16.7044
9.7424 | -34.4226 | -16.7138
9.7485 | -34.423 | -16.7234
9.7546 | -34.4234 | -16.7327
9.7607 | -34.4238 | -16.7424
9.7668 | -34.4242 | -16.7519
9.7729 | -34.4247 | -16.7614
9.779 | -34.4251 | -16.7707
9.7851 | -34.4255 | -16.7803
9.7912 | -34.426 | -16.7898
9.7973 | -34.4264 | -16.7994
9.8034 | -34.4269 | -16.8088
9.8095 | -34.4273 | -16.8184
9.8157 | -34.4277 | -16.8279
9.8218 | -34.4282 | -16.8376
9.8279 | -34.4286 | -16.847
9.834 | -34.429 | -16.8566
9.8401 | -34.4294 | -16.866
9.8462 | -34.4299 | -16.8753
9.8523 | -34.4303 | -16.8848
9.8584 | -34.4307 | -16.8943
9.8645 | -34.4311 | -16.9038
9.8706 | -34.4315 | -16.9133
9.8767 | -34.4319 | -16.9227
9.8828 | -34.4323 | -16.9321
9.8889 | -34.4327 | -16.9414
9.895 | -34.4332 | -16.9509
9.9011 | -34.4336 | -16.9604
9.9072 | -34.434 | -16.9699
9.9133 | -34.4344 | -16.9793
9.9194 | -34.4349 | -16.9887
9.9255 | -34.4353 | -16.9981
9.9316 | -34.4357 | -17.0075
9.9377 | -34.4361 | -17.0169
9.9438 | -34.4365 | -17.0265
9.9499 | -34.437 | -17.0361
9.956 | -34.4374 | -17.0456
9.9622 | -34.4378 | -17.0551
9.9683 | -34.4382 | -17.0645
9.9744 | -34.4386 | -17.0742
9.9805 | -34.4391 | -17.0836
9.9866 | -34.4395 | -17.093
9.9927 | -34.4399 | -17.1024
9.9988 | -34.4403 | -17.1119
10.0049 | -34.4408 | -17.1213
10.011 | -34.4412 | -17.1309
10.0171 | -34.4416 | -17.1402
10.0232 | -34.4421 | -17.1496
10.0293 | -34.4425 | -17.1591
10.0354 | -34.443 | -17.1685
10.0415 | -34.4434 | -17.178
10.0476 | -34.4438 | -17.1875
10.0537 | -34.4442 | -17.1971
10.0598 | -34.4447 | -17.2065
10.0659 | -34.4451 | -17.216
10.072 | -34.4455 | -17.2253
10.0781 | -34.4459 | -17.2347
10.0842 | -34.4463 | -17.2441
10.0903 | -34.4468 | -17.2535
10.0964 | -34.4472 | -17.2627
10.1026 | -34.4476 | -17.2722
10.1087 | -34.448 | -17.2817
10.1148 | -34.4484 | -17.2912
10.1209 | -34.4489 | -17.3005
10.127 | -34.4493 | -17.3097
10.1331 | -34.4497 | -17.3191
10.1392 | -34.4501 | -17.3283
10.1453 | -34.4506 | -17.3377
10.1514 | -34.451 | -17.3472
10.1575 | -34.4514 | -17.3567
10.1636 | -34.4518 | -17.3659
10.1697 | -34.4523 | -17.3754
10.1758 | -34.4527 | -17.3846
10.1819 | -34.4531 | -17.394
10.188 | -34.4535 | -17.4034
10.1941 | -34.4539 | -17.4128
10.2002 | -34.4543 | -17.4223
10.2063 | -34.4548 | -17.4317
10.2124 | -34.4552 | -17.4411
10.2185 | -34.4556 | -17.4506
10.2246 | -34.456 | -17.4599
10.2307 | -34.4565 | -17.4694
10.2368 | -34.4569 | -17.4785
10.2429 | -34.4573 | -17.4879
10.2491 | -34.4577 | -17.4974
10.2552 | -34.4581 | -17.5068
10.2613 | -34.4585 | -17.5162
10.2674 | -34.459 | -17.5257
10.2735 | -34.4594 | -17.5352
10.2796 | -34.4598 | -17.5446
10.2857 | -34.4603 | -17.5541
10.2918 | -34.4607 | -17.5636
10.2979 | -34.4611 | -17.5729
10.304 | -34.4616 | -17.5821
10.3101 | -34.462 | -17.5916
10.3162 | -34.4624 | -17.6011
10.3223 | -34.4628 | -17.6103
10.3284 | -34.4632 | -17.6197
10.3345 | -34.4636 | -17.629
10.3406 | -34.464 | -17.6383
10.3467 | -34.4644 | -17.6477
10.3528 | -34.4649 | -17.6569
10.3589 | -34.4653 | -17.6662
10.365 | -34.4657 | -17.6758
10.3711 | -34.4661 | -17.6852
10.3772 | -34.4666 | -17.6945
10.3833 | -34.467 | -17.7039
10.3895 | -34.4674 | -17.7132
10.3956 | -34.4678 | -17.7226
10.4017 | -34.4683 | -17.7318
10.4078 | -34.4687 | -17.7412
10.4139 | -34.4691 | -17.7506
10.42 | -34.4695 | -17.7601
10.4261 | -34.47 | -17.7695
10.4322 | -34.4704 | -17.7788
10.4383 | -34.4708 | -17.7881
10.4444 | -34.4712 | -17.7973
10.4505 | -34.4716 | -17.8067
10.4566 | -34.472 | -17.8159
10.4627 | -34.4724 | -17.8252
10.4688 | -34.4729 | -17.8344
10.4749 | -34.4733 | -17.8438
10.481 | -34.4737 | -17.8532
10.4871 | -34.4741 | -17.8624
10.4932 | -34.4746 | -17.8716
10.4993 | -34.475 | -17.881
10.5054 | -34.4754 | -17.8904
10.5115 | -34.4758 | -17.8998
10.5176 | -34.4763 | -17.909
10.5237 | -34.4767 | -17.9186
10.5298 | -34.4771 | -17.9277
10.536 | -34.4775 | -17.9372
10.5421 | -34.4779 | -17.9464
10.5482 | -34.4783 | -17.9558
10.5543 | -34.4788 | -17.9651
10.5604 | -34.4792 | -17.9743
10.5665 | -34.4796 | -17.9834
10.5726 | -34.48 | -17.9928
10.5787 | -34.4804 | -18.0021
10.5848 | -34.4809 | -18.0111
10.5909 | -34.4813 | -18.0206
10.597 | -34.4817 | -18.03
10.6031 | -34.4821 | -18.0395
10.6092 | -34.4825 | -18.0488
10.6153 | -34.483 | -18.0583
10.6214 | -34.4834 | -18.0674
10.6275 | -34.4838 | -18.0768
10.6336 | -34.4842 | -18.0861
10.6397 | -34.4846 | -18.0955
10.6458 | -34.4851 | -18.1047
10.6519 | -34.4855 | -18.1142
10.658 | -34.4859 | -18.1235
10.6641 | -34.4863 | -18.1329
10.6702 | -34.4867 | -18.1423
10.6764 | -34.4872 | -18.1517
10.6825 | -34.4876 | -18.1609
10.6886 | -34.488 | -18.17
10.6947 | -34.4884 | -18.1794
10.7008 | -34.4889 | -18.1886
10.7069 | -34.4893 | -18.198
10.713 | -34.4897 | -18.2072
10.7191 | -34.4901 | -18.2165
10.7252 | -34.4905 | -18.2256
10.7313 | -34.4909 | -18.2349
10.7374 | -34.4913 | -18.2443
10.7435 | -34.4917 | -18.2536
10.7496 | -34.4922 | -18.2627
10.7557 | -34.4926 | -18.2721
10.7618 | -34.493 | -18.2812
10.7679 | -34.4934 | -18.2905
10.774 | -34.4938 | -18.2998
10.7801 | -34.4943 | -18.3092
10.7862 | -34.4947 | -18.3183
10.7923 | -34.4951 | -18.3278
10.7984 | -34.4955 | -18.3369
10.8045 | -34.496 | -18.346
10.8106 | -34.4964 | -18.3554
10.8168 | -34.4968 | -18.3646
10.8229 | -34.4972 | -18.3739
10.829 | -34.4976 | -18.3831
10.8351 | -34.498 | -18.3921
10.8412 | -34.4985 | -18.4013
10.8473 | -34.4989 | -18.4104
10.8534 | -34.4993 | -18.4197
10.8595 | -34.4998 | -18.4289
10.8656 | -34.5002 | -18.4379
10.8717 | -34.5006 | -18.4472
10.8778 | -34.5011 | -18.4565
10.8839 | -34.5015 | -18.4658
10.89 | -34.5019 | -18.4751
10.8961 | -34.5024 | -18.4843
10.9022 | -34.5028 | -18.4935
10.9083 | -34.5032 | -18.5027
10.9144 | -34.5036 | -18.5119
10.9205 | -34.504 | -18.5211
10.9266 | -34.5044 | -18.5303
10.9327 | -34.5049 | -18.5394
10.9388 | -34.5053 | -18.5485
10.9449 | -34.5057 | -18.5577
10.951 | -34.5061 | -18.5669
10.9571 | -34.5066 | -18.5762
10.9633 | -34.507 | -18.5855
10.9694 | -34.5074 | -18.5949
10.9755 | -34.5078 | -18.6042
10.9816 | -34.5082 | -18.6135
10.9877 | -34.5087 | -18.6226
10.9938 | -34.5091 | -18.6319
10.9999 | -34.5096 | -18.641
11.006 | -34.51 | -18.65
11.0121 | -34.5104 | -18.6593
11.0182 | -34.5108 | -18.6685
11.0243 | -34.5113 | -18.6777
11.0304 | -34.5117 | -18.6869
11.0365 | -34.5121 | -18.696
11.0426 | -34.5126 | -18.7051
11.0487 | -34.513 | -18.7144
11.0548 | -34.5134 | -18.7235
11.0609 | -34.5139 | -18.7328
11.067 | -34.5143 | -18.742
11.0731 | -34.5147 | -18.7511
11.0792 | -34.5152 | -18.7602
11.0853 | -34.5156 | -18.7694
11.0914 | -34.516 | -18.7787
11.0975 | -34.5164 | -18.7879
11.1037 | -34.5168 | -18.797
11.1098 | -34.5173 | -18.8063
11.1159 | -34.5177 | -18.8152
11.122 | -34.5181 | -18.8243
11.1281 | -34.5185 | -18.8335
11.1342 | -34.519 | -18.8427
11.1403 | -34.5194 | -18.852
11.1464 | -34.5198 | -18.8612
11.1525 | -34.5203 | -18.8703
11.1586 | -34.5207 | -18.8796
11.1647 | -34.5211 | -18.8886
11.1708 | -34.5215 | -18.8978
11.1769 | -34.522 | -18.907
11.183 | -34.5224 | -18.9161
11.1891 | -34.5229 | -18.9252
11.1952 | -34.5233 | -18.9343
11.2013 | -34.5237 | -18.9435
11.2074 | -34.5242 | -18.9527
11.2135 | -34.5246 | -18.962
11.2196 | -34.525 | -18.9712
11.2257 | -34.5255 | -18.9804
11.2318 | -34.5259 | -18.9895
11.2379 | -34.5264 | -18.9988
11.244 | -34.5268 | -19.0081
11.2502 | -34.5273 | -19.0172
11.2563 | -34.5277 | -19.0264
11.2624 | -34.5282 | -19.0355
11.2685 | -34.5286 | -19.0446
11.2746 | -34.529 | -19.0537
11.2807 | -34.5294 | -19.063
11.2868 | -34.5299 | -19.0723
11.2929 | -34.5303 | -19.0811
11.299 | -34.5308 | -19.0903
11.3051 | -34.5312 | -19.0996
11.3112 | -34.5317 | -19.1088
11.3173 | -34.5321 | -19.1178
11.3234 | -34.5326 | -19.1268
11.3295 | -34.533 | -19.136
11.3356 | -34.5334 | -19.1452
11.3417 | -34.5339 | -19.1542
11.3478 | -34.5343 | -19.1636
11.3539 | -34.5347 | -19.1727
11.36 | -34.5351 | -19.1817
11.3661 | -34.5356 | -19.1907
11.3722 | -34.536 | -19.1996
11.3783 | -34.5365 | -19.2088
11.3844 | -34.5369 | -19.218
11.3906 | -34.5374 | -19.2271
11.3967 | -34.5378 | -19.2361
11.4028 | -34.5382 | -19.2454
11.4089 | -34.5387 | -19.2544
11.415 | -34.5391 | -19.2635
11.4211 | -34.5395 | -19.2726
11.4272 | -34.54 | -19.2818
11.4333 | -34.5404 | -19.2906
11.4394 | -34.5408 | -19.2997
11.4455 | -34.5412 | -19.3088
11.4516 | -34.5417 | -19.3179
11.4577 | -34.5421 | -19.327
11.4638 | -34.5426 | -19.336
11.4699 | -34.543 | -19.3452
11.476 | -34.5435 | -19.3544
11.4821 | -34.5439 | -19.3635
11.4882 | -34.5444 | -19.3727
11.4943 | -34.5448 | -19.3819
11.5004 | -34.5453 | -19.3908
11.5065 | -34.5457 | -19.3998
11.5126 | -34.5462 | -19.4088
11.5187 | -34.5466 | -19.418
11.5248 | -34.547 | -19.4273
11.5309 | -34.5475 | -19.4363
11.5371 | -34.5479 | -19.4453
11.5432 | -34.5484 | -19.4544
11.5493 | -34.5488 | -19.4637
11.5554 | -34.5493 | -19.4728
11.5615 | -34.5497 | -19.482
11.5676 | -34.5502 | -19.4911
11.5737 | -34.5506 | -19.5001
11.5798 | -34.5511 | -19.509
11.5859 | -34.5515 | -19.5181
11.592 | -34.5519 | -19.5271
11.5981 | -34.5524 | -19.5361
11.6042 | -34.5529 | -19.5452
11.6103 | -34.5533 | -19.5542
11.6164 | -34.5537 | -19.5631
11.6225 | -34.5542 | -19.5721
11.6286 | -34.5546 | -19.5812
11.6347 | -34.555 | -19.5903
11.6408 | -34.5555 | -19.5993
11.6469 | -34.5559 | -19.6083
11.653 | -34.5564 | -19.6173
11.6591 | -34.5568 | -19.6264
11.6652 | -34.5572 | -19.6354
11.6713 | -34.5577 | -19.6443
11.6775 | -34.5582 | -19.6532
11.6836 | -34.5586 | -19.6623
11.6897 | -34.5591 | -19.6713
11.6958 | -34.5595 | -19.6805
11.7019 | -34.56 | -19.6896
11.708 | -34.5604 | -19.6986
11.7141 | -34.5609 | -19.7074
11.7202 | -34.5613 | -19.7164
11.7263 | -34.5618 | -19.7255
11.7324 | -34.5622 | -19.7347
11.7385 | -34.5626 | -19.7438
11.7446 | -34.5631 | -19.7529
11.7507 | -34.5635 | -19.7617
11.7568 | -34.564 | -19.7708
11.7629 | -34.5644 | -19.7799
11.769 | -34.5649 | -19.7888
11.7751 | -34.5653 | -19.7978
11.7812 | -34.5658 | -19.8066
11.7873 | -34.5662 | -19.8156
11.7934 | -34.5667 | -19.8247
11.7995 | -34.5671 | -19.8337
11.8056 | -34.5676 | -19.8426
11.8117 | -34.568 | -19.8516
11.8178 | -34.5685 | -19.8605
11.824 | -34.569 | -19.8694
11.8301 | -34.5694 | -19.8783
11.8362 | -34.5699 | -19.8872
11.8423 | -34.5703 | -19.8962
11.8484 | -34.5707 | -19.9052
11.8545 | -34.5712 | -19.9141
11.8606 | -34.5717 | -19.9229
11.8667 | -34.5721 | -19.932
11.8728 | -34.5726 | -19.9408
11.8789 | -34.573 | -19.9498
11.885 | -34.5735 | -19.9587
11.8911 | -34.5739 | -19.9676
11.8972 | -34.5744 | -19.9765
11.9033 | -34.5748 | -19.9854
11.9094 | -34.5753 | -19.9943
11.9155 | -34.5757 | -20.0032
11.9216 | -34.5762 | -20.012
11.9277 | -34.5767 | -20.0209
11.9338 | -34.5771 | -20.0298
11.9399 | -34.5775 | -20.0388
11.946 | -34.578 | -20.0477
11.9521 | -34.5784 | -20.0565
11.9582 | -34.5789 | -20.0654
11.9644 | -34.5794 | -20.0742
11.9705 | -34.5798 | -20.0831
11.9766 | -34.5803 | -20.092
11.9827 | -34.5808 | -20.1009
11.9888 | -34.5813 | -20.1097
11.9949 | -34.5817 | -20.1183
12.001 | -34.5822 | -20.1274
12.0071 | -34.5826 | -20.1362
12.0132 | -34.5831 | -20.145
12.0193 | -34.5835 | -20.154
12.0254 | -34.584 | -20.1629
12.0315 | -34.5844 | -20.1715
12.0376 | -34.5849 | -20.1803
12.0437 | -34.5853 | -20.1891
12.0498 | -34.5858 | -20.198
12.0559 | -34.5862 | -20.2068
12.062 | -34.5867 | -20.2154
12.0681 | -34.5872 | -20.2244
12.0742 | -34.5876 | -20.2332
12.0803 | -34.5881 | -20.2422
12.0864 | -34.5885 | -20.2511
12.0925 | -34.589 | -20.26
12.0986 | -34.5895 | -20.2687
12.1047 | -34.5899 | -20.2776
12.1109 | -34.5904 | -20.2863
12.117 | -34.5909 | -20.2951
12.1231 | -34.5913 | -20.3039
12.1292 | -34.5918 | -20.3127
12.1353 | -34.5923 | -20.3215
12.1414 | -34.5927 | -20.3304
12.1475 | -34.5932 | -20.3393
12.1536 | -34.5936 | -20.3482
12.1597 | -34.5941 | -20.357
12.1658 | -34.5946 | -20.3657
12.1719 | -34.595 | -20.3743
12.178 | -34.5955 | -20.3832
12.1841 | -34.5959 | -20.3922
12.1902 | -34.5964 | -20.401
12.1963 | -34.5969 | -20.4099
12.2024 | -34.5974 | -20.4184
12.2085 | -34.5978 | -20.4272
12.2146 | -34.5983 | -20.4361
12.2207 | -34.5988 | -20.4448
12.2268 | -34.5992 | -20.4536
12.2329 | -34.5997 | -20.4624
12.239 | -34.6002 | -20.4708
12.2451 | -34.6007 | -20.4796
12.2513 | -34.6011 | -20.4885
12.2574 | -34.6016 | -20.4973
12.2635 | -34.602 | -20.5061
12.2696 | -34.6025 | -20.5148
12.2757 | -34.603 | -20.5233
12.2818 | -34.6035 | -20.5321
12.2879 | -34.6039 | -20.5409
12.294 | -34.6044 | -20.5495
12.3001 | -34.6049 | -20.5583
12.3062 | -34.6053 | -20.5671
12.3123 | -34.6058 | -20.5755
12.3184 | -34.6063 | -20.5842
12.3245 | -34.6067 | -20.5928
12.3306 | -34.6072 | -20.6016
12.3367 | -34.6077 | -20.6103
12.3428 | -34.6081 | -20.619
12.3489 | -34.6086 | -20.6275
12.355 | -34.6091 | -20.6364
12.3611 | -34.6096 | -20.6452
12.3672 | -34.61 | -20.6539
12.3733 | -34.6105 | -20.6626
12.3794 | -34.6109 | -20.671
12.3855 | -34.6114 | -20.6796
12.3916 | -34.6119 | -20.6883
12.3978 | -34.6123 | -20.6972
12.4039 | -34.6128 | -20.706
12.41 | -34.6133 | -20.7147
12.4161 | -34.6138 | -20.7231
12.4222 | -34.6142 | -20.7317
12.4283 | -34.6147 | -20.7404
12.4344 | -34.6152 | -20.7492
12.4405 | -34.6156 | -20.7579
12.4466 | -34.6161 | -20.7665
12.4527 | -34.6166 | -20.7751
12.4588 | -34.617 | -20.7838
12.4649 | -34.6175 | -20.7925
12.471 | -34.618 | -20.8013
12.4771 | -34.6184 | -20.81
12.4832 | -34.6189 | -20.8186
12.4893 | -34.6194 | -20.8271
12.4954 | -34.6198 | -20.8357
12.5015 | -34.6203 | -20.8443
12.5076 | -34.6208 | -20.853
12.5137 | -34.6212 | -20.8617
12.5198 | -34.6217 | -20.8702
12.5259 | -34.6222 | -20.8789
12.532 | -34.6226 | -20.8876
12.5382 | -34.6231 | -20.8963
12.5443 | -34.6236 | -20.9049
12.5504 | -34.624 | -20.9136
12.5565 | -34.6245 | -20.922
12.5626 | -34.625 | -20.9308
12.5687 | -34.6255 | -20.9395
12.5748 | -34.6259 | -20.9483
12.5809 | -34.6264 | -20.9569
12.587 | -34.6269 | -20.9656
12.5931 | -34.6274 | -20.974
12.5992 | -34.6279 | -20.9826
12.6053 | -34.6283 | -20.9913
12.6114 | -34.6288 | -20.9999
12.6175 | -34.6293 | -21.0086
12.6236 | -34.6297 | -21.0173
12.6297 | -34.6302 | -21.0258
12.6358 | -34.6307 | -21.0342
12.6419 | -34.6312 | -21.0427
12.648 | -34.6316 | -21.0514
12.6541 | -34.6321 | -21.06
12.6602 | -34.6326 | -21.0682
12.6663 | -34.633 | -21.0769
12.6724 | -34.6335 | -21.0854
12.6785 | -34.634 | -21.0939
12.6847 | -34.6345 | -21.1024
12.6908 | -34.635 | -21.111
12.6969 | -34.6355 | -21.1194
12.703 | -34.6359 | -21.1281
12.7091 | -34.6363 | -21.1367
12.7152 | -34.6368 | -21.1451
12.7213 | -34.6373 | -21.1538
12.7274 | -34.6377 | -21.1624
12.7335 | -34.6382 | -21.1707
12.7396 | -34.6387 | -21.1791
12.7457 | -34.6392 | -21.1879
12.7518 | -34.6396 | -21.1964
12.7579 | -34.6401 | -21.2051
12.764 | -34.6406 | -21.2138
12.7701 | -34.641 | -21.2224
12.7762 | -34.6415 | -21.2308
12.7823 | -34.642 | -21.2394
12.7884 | -34.6425 | -21.248
12.7945 | -34.643 | -21.2567
12.8006 | -34.6435 | -21.2651
12.8067 | -34.6439 | -21.2735
12.8128 | -34.6444 | -21.282
12.8189 | -34.6448 | -21.2905
12.8251 | -34.6453 | -21.2991
12.8312 | -34.6457 | -21.3077
12.8373 | -34.6462 | -21.3162
12.8434 | -34.6467 | -21.3248
12.8495 | -34.6472 | -21.3332
12.8556 | -34.6476 | -21.3417
12.8617 | -34.6481 | -21.3503
12.8678 | -34.6486 | -21.359
12.8739 | -34.649 | -21.3674
12.88 | -34.6495 | -21.3759
12.8861 | -34.6499 | -21.3844
12.8922 | -34.6504 | -21.3929
12.8983 | -34.6509 | -21.4014
12.9044 | -34.6514 | -21.4097
12.9105 | -34.6518 | -21.4182
12.9166 | -34.6523 | -21.4266
12.9227 | -34.6528 | -21.4351
12.9288 | -34.6532 | -21.4437
12.9349 | -34.6537 | -21.4522
12.941 | -34.6542 | -21.4606
12.9471 | -34.6547 | -21.4691
12.9532 | -34.6552 | -21.4776
12.9593 | -34.6556 | -21.4861
12.9654 | -34.6561 | -21.4947
12.9716 | -34.6566 | -21.5033
12.9777 | -34.6571 | -21.5118
12.9838 | -34.6576 | -21.5202
12.9899 | -34.658 | -21.5288
12.996 | -34.6585 | -21.5374
13.0021 | -34.659 | -21.5459
13.0082 | -34.6595 | -21.5543
13.0143 | -34.6599 | -21.5625
13.0204 | -34.6604 | -21.5709
13.0265 | -34.6609 | -21.5793
13.0326 | -34.6614 | -21.5879
13.0387 | -34.6619 | -21.5964
13.0448 | -34.6623 | -21.605
13.0509 | -34.6628 | -21.6133
13.057 | -34.6633 | -21.6217
13.0631 | -34.6638 | -21.6301
13.0692 | -34.6642 | -21.6386
13.0753 | -34.6647 | -21.6471
13.0814 | -34.6652 | -21.6558
13.0875 | -34.6657 | -21.6641
13.0936 | -34.6661 | -21.6726
13.0997 | -34.6666 | -21.6809
13.1058 | -34.6671 | -21.6893
13.112 | -34.6676 | -21.6977
13.1181 | -34.6681 | -21.7061
13.1242 | -34.6685 | -21.7146
13.1303 | -34.669 | -21.7231
13.1364 | -34.6694 | -21.7316
13.1425 | -34.6699 | -21.74
13.1486 | -34.6704 | -21.7484
13.1547 | -34.6709 | -21.7569
13.1608 | -34.6713 | -21.7654
13.1669 | -34.6718 | -21.7738
13.173 | -34.6723 | -21.7822
13.1791 | -34.6728 | -21.7906
13.1852 | -34.6733 | -21.7991
13.1913 | -34.6737 | -21.8075
13.1974 | -34.6742 | -21.816
13.2035 | -34.6747 | -21.8244
13.2096 | -34.6751 | -21.8327
13.2157 | -34.6756 | -21.8411
13.2218 | -34.6761 | -21.8495
13.2279 | -34.6766 | -21.8578
13.234 | -34.6771 | -21.8663
13.2401 | -34.6776 | -21.8747
13.2462 | -34.6781 | -21.8831
13.2524 | -34.6786 | -21.8915
13.2585 | -34.679 | -21.8998
13.2646 | -34.6795 | -21.9083
13.2707 | -34.68 | -21.9166
13.2768 | -34.6805 | -21.9251
13.2829 | -34.681 | -21.9337
13.289 | -34.6814 | -21.942
13.2951 | -34.6819 | -21.9503
13.3012 | -34.6824 | -21.9587
13.3073 | -34.6829 | -21.9672
13.3134 | -34.6833 | -21.9756
13.3195 | -34.6838 | -21.9839
13.3256 | -34.6842 | -21.9922
13.3317 | -34.6848 | -22.0004
13.3378 | -34.6852 | -22.0088
13.3439 | -34.6857 | -22.0174
13.35 | -34.6862 | -22.0259
13.3561 | -34.6866 | -22.0341
13.3622 | -34.6871 | -22.0424
13.3683 | -34.6876 | -22.0508
13.3744 | -34.6881 | -22.0592
13.3805 | -34.6885 | -22.0676
13.3866 | -34.689 | -22.0758
13.3927 | -34.6894 | -22.0844
13.3989 | -34.6899 | -22.0928
13.405 | -34.6904 | -22.1013
13.4111 | -34.6909 | -22.1095
13.4172 | -34.6913 | -22.1179
13.4233 | -34.6918 | -22.1264
13.4294 | -34.6923 | -22.1351
13.4355 | -34.6927 | -22.1432
13.4416 | -34.6932 | -22.1518
13.4477 | -34.6937 | -22.1601
13.4538 | -34.6942 | -22.1687
13.4599 | -34.6947 | -22.1772
13.466 | -34.6951 | -22.1855
13.4721 | -34.6956 | -22.1938
13.4782 | -34.6961 | -22.2021
13.4843 | -34.6966 | -22.2106
13.4904 | -34.6971 | -22.2189
13.4965 | -34.6975 | -22.2274
13.5026 | -34.698 | -22.2357
13.5087 | -34.6985 | -22.244
13.5148 | -34.6989 | -22.2525
13.5209 | -34.6994 | -22.261
13.527 | -34.6999 | -22.2694
13.5331 | -34.7004 | -22.2778
13.5393 | -34.7008 | -22.2861
13.5454 | -34.7013 | -22.2944
13.5515 | -34.7018 | -22.3031
13.5576 | -34.7023 | -22.3114
13.5637 | -34.7027 | -22.3197
13.5698 | -34.7032 | -22.3281
13.5759 | -34.7036 | -22.3362
13.582 | -34.7041 | -22.3445
13.5881 | -34.7046 | -22.3528
13.5942 | -34.7051 | -22.3611
13.6003 | -34.7056 | -22.3694
13.6064 | -34.706 | -22.3777
13.6125 | -34.7065 | -22.3863
13.6186 | -34.707 | -22.3946
13.6247 | -34.7075 | -22.4031
13.6308 | -34.708 | -22.4114
13.6369 | -34.7084 | -22.42
13.643 | -34.7089 | -22.4284
13.6491 | -34.7094 | -22.4365
13.6552 | -34.7098 | -22.4449
13.6613 | -34.7103 | -22.4533
13.6674 | -34.7108 | -22.4617
13.6735 | -34.7113 | -22.4701
13.6796 | -34.7117 | -22.4784
13.6858 | -34.7122 | -22.4865
13.6919 | -34.7127 | -22.4948
13.698 | -34.7132 | -22.5032
13.7041 | -34.7136 | -22.5118
13.7102 | -34.7141 | -22.5201
13.7163 | -34.7146 | -22.5285
13.7224 | -34.7151 | -22.5369
13.7285 | -34.7156 | -22.5452
13.7346 | -34.7161 | -22.5535
13.7407 | -34.7166 | -22.5619
13.7468 | -34.7171 | -22.5702
13.7529 | -34.7175 | -22.5784
13.759 | -34.718 | -22.5869
13.7651 | -34.7185 | -22.5954
13.7712 | -34.7189 | -22.6037
13.7773 | -34.7194 | -22.6121
13.7834 | -34.7199 | -22.6204
13.7895 | -34.7204 | -22.6288
13.7956 | -34.7209 | -22.6375
13.8017 | -34.7213 | -22.6457
13.8078 | -34.7218 | -22.654
13.8139 | -34.7223 | -22.6622
13.82 | -34.7227 | -22.6705
13.8262 | -34.7232 | -22.679
13.8323 | -34.7237 | -22.6874
13.8384 | -34.7242 | -22.6957
13.8445 | -34.7247 | -22.7042
13.8506 | -34.7251 | -22.7126
13.8567 | -34.7257 | -22.7209
13.8628 | -34.7261 | -22.7292
13.8689 | -34.7266 | -22.7376
13.875 | -34.7271 | -22.746
13.8811 | -34.7275 | -22.7544
13.8872 | -34.728 | -22.7627
13.8933 | -34.7285 | -22.7708
13.8994 | -34.729 | -22.7792
13.9055 | -34.7294 | -22.7877
13.9116 | -34.7299 | -22.7961
13.9177 | -34.7304 | -22.8044
13.9238 | -34.7309 | -22.8127
13.9299 | -34.7313 | -22.821
13.936 | -34.7318 | -22.8295
13.9421 | -34.7323 | -22.8379
13.9482 | -34.7327 | -22.846
13.9543 | -34.7332 | -22.8544
13.9604 | -34.7336 | -22.8626
13.9665 | -34.7342 | -22.8709
13.9727 | -34.7347 | -22.8793
13.9788 | -34.7352 | -22.8876
13.9849 | -34.7356 | -22.8959
13.991 | -34.7361 | -22.9042
13.9971 | -34.7366 | -22.9123
14.0032 | -34.7371 | -22.9208
14.0093 | -34.7375 | -22.9291
14.0154 | -34.738 | -22.9373
14.0215 | -34.7385 | -22.9456
14.0276 | -34.739 | -22.9539
14.0337 | -34.7395 | -22.9623
14.0398 | -34.74 | -22.9706
14.0459 | -34.7404 | -22.9789
14.052 | -34.7409 | -22.9873
14.0581 | -34.7414 | -22.9957
14.0642 | -34.7419 | -23.0039
14.0703 | -34.7423 | -23.0121
14.0764 | -34.7428 | -23.0203
14.0825 | -34.7433 | -23.0286
14.0886 | -34.7438 | -23.0371
14.0947 | -34.7442 | -23.0453
14.1008 | -34.7447 | -23.0536
14.1069 | -34.7452 | -23.0618
14.1131 | -34.7457 | -23.0703
14.1192 | -34.7461 | -23.0785
14.1253 | -34.7466 | -23.0865
14.1314 | -34.7471 | -23.0948
14.1375 | -34.7476 | -23.1029
14.1436 | -34.7481 | -23.1111
14.1497 | -34.7486 | -23.1195
14.1558 | -34.7491 | -23.1278
14.1619 | -34.7496 | -23.1361
14.168 | -34.75 | -23.1443
14.1741 | -34.7505 | -23.1525
14.1802 | -34.751 | -23.1608
14.1863 | -34.7515 | -23.1691
14.1924 | -34.752 | -23.1774
14.1985 | -34.7525 | -23.1856
14.2046 | -34.753 | -23.1941
14.2107 | -34.7534 | -23.2023
14.2168 | -34.7539 | -23.2106
14.2229 | -34.7544 | -23.2188
14.229 | -34.7549 | -23.2272
14.2351 | -34.7553 | -23.2355
14.2412 | -34.7558 | -23.2439
14.2473 | -34.7563 | -23.2521
14.2534 | -34.7568 | -23.2603
14.2596 | -34.7572 | -23.2686
14.2657 | -34.7577 | -23.2769
14.2718 | -34.7582 | -23.2853
14.2779 | -34.7587 | -23.2936
14.284 | -34.7592 | -23.3019
14.2901 | -34.7597 | -23.3101
14.2962 | -34.7602 | -23.3184
14.3023 | -34.7606 | -23.3267
14.3084 | -34.7611 | -23.3349
14.3145 | -34.7616 | -23.3432
14.3206 | -34.7621 | -23.3516
14.3267 | -34.7626 | -23.3599
14.3328 | -34.763 | -23.3682
14.3389 | -34.7635 | -23.3763
14.345 | -34.764 | -23.3846
14.3511 | -34.7644 | -23.3929
14.3572 | -34.7649 | -23.4013
14.3633 | -34.7654 | -23.4098
14.3694 | -34.7659 | -23.418
14.3755 | -34.7664 | -23.4264
14.3816 | -34.7669 | -23.4346
14.3877 | -34.7674 | -23.4427
14.3938 | -34.7679 | -23.4512
14.4 | -34.7684 | -23.4594
14.4061 | -34.7688 | -23.4677
14.4122 | -34.7693 | -23.476
14.4183 | -34.7698 | -23.4841
14.4244 | -34.7703 | -23.4925
14.4305 | -34.7708 | -23.5007
14.4366 | -34.7713 | -23.5091
14.4427 | -34.7718 | -23.5176
14.4488 | -34.7723 | -23.5258
14.4549 | -34.7727 | -23.5341
14.461 | -34.7732 | -23.5422
14.4671 | -34.7737 | -23.5506
14.4732 | -34.7742 | -23.5589
14.4793 | -34.7747 | -23.5671
14.4854 | -34.7752 | -23.5754
14.4915 | -34.7757 | -23.5834
14.4976 | -34.7762 | -23.5915
14.5037 | -34.7767 | -23.5997
14.5098 | -34.7771 | -23.608
14.5159 | -34.7776 | -23.6163
14.522 | -34.7781 | -23.6245
14.5281 | -34.7786 | -23.6326
14.5342 | -34.7791 | -23.6409
14.5403 | -34.7796 | -23.6491
14.5465 | -34.78 | -23.6573
14.5526 | -34.7805 | -23.6657
14.5587 | -34.781 | -23.674
14.5648 | -34.7815 | -23.6821
14.5709 | -34.782 | -23.6902
14.577 | -34.7825 | -23.6984
14.5831 | -34.783 | -23.7067
14.5892 | -34.7835 | -23.7149
14.5953 | -34.784 | -23.723
14.6014 | -34.7844 | -23.7312
14.6075 | -34.7849 | -23.7394
14.6136 | -34.7854 | -23.7477
14.6197 | -34.7858 | -23.756
14.6258 | -34.7863 | -23.7643
14.6319 | -34.7868 | -23.7724
14.638 | -34.7873 | -23.7806
14.6441 | -34.7878 | -23.789
14.6502 | -34.7883 | -23.7973
14.6563 | -34.7887 | -23.8056
14.6624 | -34.7892 | -23.8138
14.6685 | -34.7897 | -23.8218
14.6746 | -34.7902 | -23.8301
14.6807 | -34.7907 | -23.8382
14.6869 | -34.7912 | -23.8464
14.693 | -34.7917 | -23.8545
14.6991 | -34.7922 | -23.8627
14.7052 | -34.7927 | -23.8707
14.7113 | -34.7932 | -23.8789
14.7174 | -34.7937 | -23.8872
14.7235 | -34.7942 | -23.8954
14.7296 | -34.7946 | -23.9035
14.7357 | -34.7952 | -23.9117
14.7418 | -34.7957 | -23.9199
14.7479 | -34.7961 | -23.9281
14.754 | -34.7966 | -23.9364
14.7601 | -34.7971 | -23.9445
14.7662 | -34.7976 | -23.9527
14.7723 | -34.7981 | -23.9606
14.7784 | -34.7986 | -23.9689
14.7845 | -34.7991 | -23.9771
14.7906 | -34.7995 | -23.9854
14.7967 | -34.8 | -23.9935
14.8028 | -34.8005 | -24.0017
14.8089 | -34.801 | -24.0099
14.815 | -34.8015 | -24.0181
14.8211 | -34.802 | -24.0263
14.8272 | -34.8025 | -24.0345
14.8334 | -34.803 | -24.0426
14.8395 | -34.8035 | -24.0507
14.8456 | -34.804 | -24.0588
14.8517 | -34.8045 | -24.0671
14.8578 | -34.805 | -24.0753
14.8639 | -34.8054 | -24.0835
14.87 | -34.8059 | -24.0918
14.8761 | -34.8064 | -24.0999
14.8822 | -34.8069 | -24.1081
14.8883 | -34.8074 | -24.1163
14.8944 | -34.8079 | -24.1244
14.9005 | -34.8083 | -24.1326
14.9066 | -34.8088 | -24.1409
14.9127 | -34.8094 | -24.1489
14.9188 | -34.8099 | -24.1571
14.9249 | -34.8103 | -24.1652
14.931 | -34.8108 | -24.1734
14.9371 | -34.8113 | -24.1815
14.9432 | -34.8118 | -24.1896
14.9493 | -34.8123 | -24.1977
14.9554 | -34.8128 | -24.2059
14.9615 | -34.8133 | -24.214
14.9676 | -34.8138 | -24.2223
14.9738 | -34.8143 | -24.2303
14.9799 | -34.8148 | -24.2385
14.986 | -34.8152 | -24.2465
14.9921 | -34.8157 | -24.2546
14.9982 | -34.8162 | -24.2628
15.0043 | -34.8167 | -24.2708
15.0104 | -34.8172 | -24.279
15.0165 | -34.8177 | -24.2873
15.0226 | -34.8182 | -24.295
15.0287 | -34.8187 | -24.3032
15.0348 | -34.8192 | -24.3114
15.0409 | -34.8197 | -24.3196
15.047 | -34.8202 | -24.3278
15.0531 | -34.8206 | -24.3359
15.0592 | -34.8211 | -24.3441
15.0653 | -34.8216 | -24.3522
15.0714 | -34.8221 | -24.3602
15.0775 | -34.8226 | -24.3682
15.0836 | -34.8231 | -24.3763
15.0897 | -34.8236 | -24.3841
15.0958 | -34.8241 | -24.3922
15.1019 | -34.8246 | -24.4003
15.108 | -34.8251 | -24.4083
15.1141 | -34.8256 | -24.4163
15.1203 | -34.8261 | -24.4245
15.1264 | -34.8266 | -24.4325
15.1325 | -34.8271 | -24.4405
15.1386 | -34.8276 | -24.4484
15.1447 | -34.8281 | -24.4564
15.1508 | -34.8287 | -24.4645
15.1569 | -34.8292 | -24.4726
15.163 | -34.8297 | -24.4806
15.1691 | -34.8302 | -24.4886
15.1752 | -34.8307 | -24.4967
15.1813 | -34.8312 | -24.5048
15.1874 | -34.8317 | -24.5129
15.1935 | -34.8322 | -24.5208
15.1996 | -34.8327 | -24.5288
15.2057 | -34.8332 | -24.5369
15.2118 | -34.8337 | -24.545
15.2179 | -34.8342 | -24.5532
15.224 | -34.8346 | -24.5612
15.2301 | -34.8352 | -24.569
15.2362 | -34.8357 | -24.5772
15.2423 | -34.8362 | -24.5853
15.2484 | -34.8367 | -24.5934
15.2545 | -34.8372 | -24.6016
15.2607 | -34.8377 | -24.6097
15.2668 | -34.8382 | -24.6174
15.2729 | -34.8387 | -24.6254
15.279 | -34.8392 | -24.6336
15.2851 | -34.8397 | -24.6415
15.2912 | -34.8402 | -24.6496
15.2973 | -34.8407 | -24.6578
15.3034 | -34.8412 | -24.6657
15.3095 | -34.8417 | -24.6737
15.3156 | -34.8423 | -24.6817
15.3217 | -34.8428 | -24.6897
15.3278 | -34.8433 | -24.6979
15.3339 | -34.8438 | -24.7056
15.34 | -34.8443 | -24.7135
15.3461 | -34.8448 | -24.7216
15.3522 | -34.8453 | -24.7296
15.3583 | -34.8458 | -24.7378
15.3644 | -34.8463 | -24.7458
15.3705 | -34.8468 | -24.7535
15.3766 | -34.8473 | -24.7617
15.3827 | -34.8478 | -24.7697
15.3888 | -34.8483 | -24.7777
15.3949 | -34.8488 | -24.7858
15.401 | -34.8493 | -24.7939
15.4072 | -34.8498 | -24.8017
15.4133 | -34.8503 | -24.8097
15.4194 | -34.8508 | -24.8178
15.4255 | -34.8513 | -24.8258
15.4316 | -34.8517 | -24.8339
15.4377 | -34.8522 | -24.8419
15.4438 | -34.8528 | -24.8498
15.4499 | -34.8533 | -24.8576
15.456 | -34.8538 | -24.8657
15.4621 | -34.8543 | -24.8737
15.4682 | -34.8548 | -24.8817
15.4743 | -34.8553 | -24.8896
15.4804 | -34.8558 | -24.8976
15.4865 | -34.8562 | -24.9056
15.4926 | -34.8567 | -24.9137
15.4987 | -34.8572 | -24.9216
15.5048 | -34.8577 | -24.9296
15.5109 | -34.8582 | -24.9375
15.517 | -34.8588 | -24.9455
15.5231 | -34.8593 | -24.9535
15.5292 | -34.8597 | -24.9614
15.5353 | -34.8602 | -24.9696
15.5414 | -34.8607 | -24.9773
15.5476 | -34.8612 | -24.985
15.5537 | -34.8617 | -24.993
15.5598 | -34.8622 | -25.0011
15.5659 | -34.8627 | -25.009
15.572 | -34.8632 | -25.017
15.5781 | -34.8637 | -25.0249
15.5842 | -34.8642 | -25.0327
15.5903 | -34.8647 | -25.0405
15.5964 | -34.8652 | -25.0484
15.6025 | -34.8657 | -25.0562
15.6086 | -34.8663 | -25.0642
15.6147 | -34.8668 | -25.0722
15.6208 | -34.8673 | -25.08
15.6269 | -34.8678 | -25.0878
15.633 | -34.8683 | -25.0957
15.6391 | -34.8688 | -25.1037
15.6452 | -34.8693 | -25.1117
15.6513 | -34.8698 | -25.1194
15.6574 | -34.8703 | -25.1274
15.6635 | -34.8708 | -25.1354
15.6696 | -34.8713 | -25.1433
15.6757 | -34.8718 | -25.1513
15.6818 | -34.8723 | -25.1593
15.688 | -34.8728 | -25.167
15.6941 | -34.8733 | -25.1749
15.7002 | -34.8738 | -25.1828
15.7063 | -34.8743 | -25.1906
15.7124 | -34.8748 | -25.1987
15.7185 | -34.8753 | -25.2065
15.7246 | -34.8758 | -25.2143
15.7307 | -34.8763 | -25.2221
15.7368 | -34.8768 | -25.23
15.7429 | -34.8773 | -25.238
15.749 | -34.8778 | -25.246
15.7551 | -34.8783 | -25.254
15.7612 | -34.8788 | -25.2617
15.7673 | -34.8793 | -25.2696
15.7734 | -34.8798 | -25.2775
15.7795 | -34.8803 | -25.2856
15.7856 | -34.8808 | -25.2934
15.7917 | -34.8813 | -25.3009
15.7978 | -34.8818 | -25.309
15.8039 | -34.8823 | -25.3169
15.81 | -34.8828 | -25.3248
15.8161 | -34.8833 | -25.3326
15.8222 | -34.8838 | -25.3405
15.8283 | -34.8843 | -25.3481
15.8345 | -34.8848 | -25.3561
15.8406 | -34.8853 | -25.364
15.8467 | -34.8859 | -25.3719
15.8528 | -34.8864 | -25.3798
15.8589 | -34.8869 | -25.3877
15.865 | -34.8874 | -25.3953
15.8711 | -34.8879 | -25.4033
15.8772 | -34.8884 | -25.411
15.8833 | -34.8889 | -25.4188
15.8894 | -34.8894 | -25.4267
15.8955 | -34.8899 | -25.4346
15.9016 | -34.8904 | -25.4423
15.9077 | -34.8909 | -25.45
15.9138 | -34.8913 | -25.4581
15.9199 | -34.8918 | -25.4659
15.926 | -34.8923 | -25.4737
15.9321 | -34.8928 | -25.4813
15.9382 | -34.8933 | -25.4891
15.9443 | -34.8938 | -25.4968
15.9504 | -34.8943 | -25.5047
15.9565 | -34.8948 | -25.5125
15.9626 | -34.8953 | -25.5205
15.9687 | -34.8958 | -25.528
15.9749 | -34.8963 | -25.5357
15.981 | -34.8968 | -25.5435
15.9871 | -34.8973 | -25.5514
15.9932 | -34.8978 | -25.559
15.9993 | -34.8983 | -25.5669
16.0054 | -34.8988 | -25.5744
16.0115 | -34.8993 | -25.5822
16.0176 | -34.8998 | -25.5902
16.0237 | -34.9004 | -25.5981
16.0298 | -34.9008 | -25.6058
16.0359 | -34.9013 | -25.6137
16.042 | -34.9019 | -25.6214
16.0481 | -34.9023 | -25.6292
16.0542 | -34.9028 | -25.637
16.0603 | -34.9033 | -25.6447
16.0664 | -34.9038 | -25.6525
16.0725 | -34.9043 | -25.66
16.0786 | -34.9048 | -25.6677
16.0847 | -34.9053 | -25.6757
16.0908 | -34.9058 | -25.6835
16.0969 | -34.9063 | -25.6914
16.103 | -34.9068 | -25.6992
16.1091 | -34.9074 | -25.7068
16.1152 | -34.9078 | -25.7146
16.1214 | -34.9084 | -25.7224
16.1275 | -34.9089 | -25.7302
16.1336 | -34.9094 | -25.738
16.1397 | -34.9099 | -25.7457
16.1458 | -34.9104 | -25.7533
16.1519 | -34.9109 | -25.7612
16.158 | -34.9114 | -25.7691
16.1641 | -34.9119 | -25.7769
16.1702 | -34.9124 | -25.7848
16.1763 | -34.9129 | -25.7926
16.1824 | -34.9134 | -25.7999
16.1885 | -34.9139 | -25.8076
16.1946 | -34.9144 | -25.8152
16.2007 | -34.915 | -25.8229
16.2068 | -34.9155 | -25.8306
16.2129 | -34.916 | -25.8385
16.219 | -34.9165 | -25.846
16.2251 | -34.917 | -25.8538
16.2312 | -34.9175 | -25.8615
16.2373 | -34.9179 | -25.8692
16.2434 | -34.9184 | -25.877
16.2495 | -34.919 | -25.8845
16.2556 | -34.9195 | -25.8925
16.2618 | -34.92 | -25.9002
16.2679 | -34.9205 | -25.908
16.274 | -34.921 | -25.9159
16.2801 | -34.9214 | -25.9235
16.2862 | -34.9219 | -25.9312
16.2923 | -34.9224 | -25.939
16.2984 | -34.9229 | -25.9467
16.3045 | -34.9234 | -25.9545
16.3106 | -34.9239 | -25.9622
16.3167 | -34.9244 | -25.9699
16.3228 | -34.9249 | -25.9775
16.3289 | -34.9254 | -25.9852
16.335 | -34.9259 | -25.9928
16.3411 | -34.9264 | -26.0006
16.3472 | -34.9269 | -26.0085
16.3533 | -34.9274 | -26.0163
16.3594 | -34.9279 | -26.0237
16.3655 | -34.9284 | -26.0314
16.3716 | -34.9289 | -26.0389
16.3777 | -34.9294 | -26.0465
16.3838 | -34.9299 | -26.0542
16.3899 | -34.9304 | -26.0617
16.396 | -34.9309 | -26.0694
16.4021 | -34.9313 | -26.077
16.4083 | -34.9318 | -26.0846
16.4144 | -34.9323 | -26.0923
16.4205 | -34.9328 | -26.1
16.4266 | -34.9333 | -26.1073
16.4327 | -34.9338 | -26.115
16.4388 | -34.9343 | -26.1227
16.4449 | -34.9348 | -26.1303
16.451 | -34.9354 | -26.1381
16.4571 | -34.9358 | -26.1459
16.4632 | -34.9363 | -26.1532
16.4693 | -34.9368 | -26.161
16.4754 | -34.9373 | -26.1687
16.4815 | -34.9378 | -26.1763
16.4876 | -34.9383 | -26.1838
16.4937 | -34.9388 | -26.1915
16.4998 | -34.9393 | -26.1989
16.5059 | -34.9398 | -26.2065
16.512 | -34.9403 | -26.2143
16.5181 | -34.9408 | -26.2222
16.5242 | -34.9413 | -26.2299
16.5303 | -34.9418 | -26.2375
16.5364 | -34.9423 | -26.245
16.5425 | -34.9428 | -26.2527
16.5487 | -34.9433 | -26.2602
16.5548 | -34.9438 | -26.2679
16.5609 | -34.9443 | -26.2755
16.567 | -34.9448 | -26.283
16.5731 | -34.9453 | -26.2907
16.5792 | -34.9458 | -26.2984
16.5853 | -34.9463 | -26.3062
16.5914 | -34.9468 | -26.3138
16.5975 | -34.9473 | -26.3215
16.6036 | -34.9478 | -26.3291
16.6097 | -34.9483 | -26.3367
16.6158 | -34.9487 | -26.3443
16.6219 | -34.9492 | -26.352
16.628 | -34.9497 | -26.3597
16.6341 | -34.9502 | -26.3672
16.6402 | -34.9507 | -26.3747
16.6463 | -34.9512 | -26.3825
16.6524 | -34.9517 | -26.3901
16.6585 | -34.9522 | -26.3978
16.6646 | -34.9527 | -26.4054
16.6707 | -34.9532 | -26.4131
16.6768 | -34.9537 | -26.4207
16.6829 | -34.9542 | -26.4283
16.689 | -34.9547 | -26.4359
16.6952 | -34.9552 | -26.4436
16.7013 | -34.9557 | -26.4512
16.7074 | -34.9562 | -26.4585
16.7135 | -34.9567 | -26.466
16.7196 | -34.9572 | -26.4736
16.7257 | -34.9577 | -26.4812
16.7318 | -34.9582 | -26.4887
16.7379 | -34.9587 | -26.4961
16.744 | -34.9592 | -26.5036
16.7501 | -34.9597 | -26.5112
16.7562 | -34.9602 | -26.5188
16.7623 | -34.9607 | -26.5262
16.7684 | -34.9612 | -26.5339
16.7745 | -34.9617 | -26.5415
16.7806 | -34.9622 | -26.5488
16.7867 | -34.9627 | -26.5564
16.7928 | -34.9632 | -26.564
16.7989 | -34.9637 | -26.5715
16.805 | -34.9642 | -26.5791
16.8111 | -34.9647 | -26.5867
16.8172 | -34.9652 | -26.5942
16.8233 | -34.9657 | -26.602
16.8294 | -34.9662 | -26.6095
16.8356 | -34.9667 | -26.6172
16.8417 | -34.9672 | -26.6247
16.8478 | -34.9677 | -26.6323
16.8539 | -34.9682 | -26.6398
16.86 | -34.9687 | -26.6474
16.8661 | -34.9692 | -26.655
16.8722 | -34.9697 | -26.6628
16.8783 | -34.9702 | -26.6704
16.8844 | -34.9707 | -26.6778
16.8905 | -34.9712 | -26.6854
16.8966 | -34.9717 | -26.693
16.9027 | -34.9722 | -26.7006
16.9088 | -34.9727 | -26.7083
16.9149 | -34.9732 | -26.7159
16.921 | -34.9737 | -26.7232
16.9271 | -34.9742 | -26.731
16.9332 | -34.9747 | -26.7384
16.9393 | -34.9753 | -26.7462
16.9454 | -34.9757 | -26.7536
16.9515 | -34.9762 | -26.7612
16.9576 | -34.9767 | -26.7684
16.9637 | -34.9772 | -26.776
16.9698 | -34.9777 | -26.7834
16.9759 | -34.9782 | -26.791
16.9821 | -34.9787 | -26.7988
16.9882 | -34.9792 | -26.8061
16.9943 | -34.9797 | -26.8136
17.0004 | -34.9802 | -26.8212
17.0065 | -34.9807 | -26.8288
17.0126 | -34.9812 | -26.8363
17.0187 | -34.9817 | -26.8441
17.0248 | -34.9822 | -26.8515
17.0309 | -34.9826 | -26.8589
17.037 | -34.9831 | -26.8664
17.0431 | -34.9836 | -26.874
17.0492 | -34.9841 | -26.8814
17.0553 | -34.9846 | -26.8889
17.0614 | -34.9851 | -26.8965
17.0675 | -34.9856 | -26.904
17.0736 | -34.9861 | -26.9116
17.0797 | -34.9865 | -26.919
17.0858 | -34.987 | -26.9266
17.0919 | -34.9875 | -26.9341
17.098 | -34.9881 | -26.9414
17.1041 | -34.9885 | -26.9489
17.1102 | -34.9891 | -26.9563
17.1163 | -34.9896 | -26.9639
17.1225 | -34.9901 | -26.9713
17.1286 | -34.9906 | -26.9786
17.1347 | -34.9911 | -26.9861
17.1408 | -34.9916 | -26.9937
17.1469 | -34.9921 | -27.0012
17.153 | -34.9926 | -27.0087
17.1591 | -34.9931 | -27.0164
17.1652 | -34.9935 | -27.0237
17.1713 | -34.9941 | -27.0311
17.1774 | -34.9946 | -27.0386
17.1835 | -34.995 | -27.0459
17.1896 | -34.9955 | -27.0533
17.1957 | -34.996 | -27.0608
17.2018 | -34.9965 | -27.0681
17.2079 | -34.997 | -27.0755
17.214 | -34.9975 | -27.0831
17.2201 | -34.998 | -27.0907
17.2262 | -34.9985 | -27.0982
17.2323 | -34.999 | -27.1058
17.2384 | -34.9995 | -27.1133
17.2445 | -35 | -27.1206
17.2506 | -35.0005 | -27.1281
17.2567 | -35.001 | -27.1356
17.2628 | -35.0015 | -27.143
17.269 | -35.002 | -27.1505
17.2751 | -35.0025 | -27.158
17.2812 | -35.0029 | -27.1656
17.2873 | -35.0034 | -27.1732
17.2934 | -35.0039 | -27.1808
17.2995 | -35.0044 | -27.1884
17.3056 | -35.0049 | -27.1958
17.3117 | -35.0054 | -27.2032
17.3178 | -35.0059 | -27.2106
17.3239 | -35.0064 | -27.218
17.33 | -35.0068 | -27.2255
17.3361 | -35.0073 | -27.2329
17.3422 | -35.0078 | -27.2401
17.3483 | -35.0083 | -27.2478
17.3544 | -35.0088 | -27.2553
17.3605 | -35.0093 | -27.2628
17.3666 | -35.0098 | -27.2703
17.3727 | -35.0102 | -27.2778
17.3788 | -35.0107 | -27.2851
17.3849 | -35.0112 | -27.2926
17.391 | -35.0117 | -27.3
17.3971 | -35.0122 | -27.3074
17.4032 | -35.0126 | -27.315
17.4094 | -35.0131 | -27.3223
17.4155 | -35.0136 | -27.3296
17.4216 | -35.0141 | -27.3368
17.4277 | -35.0146 | -27.3444
17.4338 | -35.0151 | -27.3519
17.4399 | -35.0157 | -27.3593
17.446 | -35.0161 | -27.3665
17.4521 | -35.0166 | -27.3738
17.4582 | -35.0171 | -27.3812
17.4643 | -35.0176 | -27.3885
17.4704 | -35.0181 | -27.3961
17.4765 | -35.0186 | -27.4037
17.4826 | -35.0191 | -27.411
17.4887 | -35.0196 | -27.4186
17.4948 | -35.0201 | -27.4263
17.5009 | -35.0206 | -27.4337
17.507 | -35.021 | -27.4411
17.5131 | -35.0215 | -27.4485
17.5192 | -35.022 | -27.4556
17.5253 | -35.0225 | -27.4631
17.5314 | -35.023 | -27.4706
17.5375 | -35.0234 | -27.4781
17.5436 | -35.0239 | -27.4856
17.5497 | -35.0244 | -27.4929
17.5559 | -35.0249 | -27.5001
17.562 | -35.0254 | -27.5074
17.5681 | -35.0259 | -27.5148
17.5742 | -35.0264 | -27.5224
17.5803 | -35.0269 | -27.5298
17.5864 | -35.0274 | -27.5372
17.5925 | -35.0279 | -27.5446
17.5986 | -35.0284 | -27.5519
17.6047 | -35.0289 | -27.5593
17.6108 | -35.0294 | -27.5666
17.6169 | -35.0299 | -27.5739
17.623 | -35.0304 | -27.5813
17.6291 | -35.0309 | -27.5886
17.6352 | -35.0314 | -27.596
17.6413 | -35.0319 | -27.6034
17.6474 | -35.0323 | -27.6109
17.6535 | -35.0328 | -27.6183
17.6596 | -35.0333 | -27.6256
17.6657 | -35.0338 | -27.633
17.6718 | -35.0343 | -27.6404
17.6779 | -35.0348 | -27.6477
17.684 | -35.0353 | -27.6552
17.6901 | -35.0358 | -27.6626
17.6963 | -35.0362 | -27.6697
17.7024 | -35.0367 | -27.6771
17.7085 | -35.0372 | -27.6844
17.7146 | -35.0377 | -27.6917
17.7207 | -35.0382 | -27.699
17.7268 | -35.0387 | -27.7063
17.7329 | -35.0392 | -27.7136
17.739 | -35.0397 | -27.721
17.7451 | -35.0402 | -27.7284
17.7512 | -35.0406 | -27.7356
17.7573 | -35.0411 | -27.7431
17.7634 | -35.0416 | -27.7502
17.7695 | -35.0421 | -27.7575
17.7756 | -35.0426 | -27.7649
17.7817 | -35.0431 | -27.7723
17.7878 | -35.0436 | -27.7797
17.7939 | -35.0441 | -27.787
17.8 | -35.0445 | -27.7943
17.8061 | -35.045 | -27.8016
17.8122 | -35.0456 | -27.809
17.8183 | -35.0461 | -27.8163
17.8244 | -35.0465 | -27.8236
17.8305 | -35.047 | -27.831
17.8366 | -35.0475 | -27.8381
17.8428 | -35.048 | -27.8455
17.8489 | -35.0485 | -27.8528
17.855 | -35.049 | -27.8603
17.8611 | -35.0494 | -27.8676
17.8672 | -35.0499 | -27.8749
17.8733 | -35.0504 | -27.8823
17.8794 | -35.0509 | -27.8897
17.8855 | -35.0514 | -27.8971
17.8916 | -35.0519 | -27.9045
17.8977 | -35.0523 | -27.9119
17.9038 | -35.0528 | -27.919
17.9099 | -35.0533 | -27.9263
17.916 | -35.0538 | -27.9336
17.9221 | -35.0543 | -27.9409
17.9282 | -35.0548 | -27.9483
17.9343 | -35.0553 | -27.9557
17.9404 | -35.0557 | -27.9629
17.9465 | -35.0562 | -27.9702
17.9526 | -35.0567 | -27.9775
17.9587 | -35.0571 | -27.9848
17.9648 | -35.0576 | -27.9921
17.9709 | -35.0581 | -27.9994
17.977 | -35.0586 | -28.0066
17.9832 | -35.0591 | -28.014
17.9893 | -35.0596 | -28.0214
17.9954 | -35.0601 | -28.0287
18.0015 | -35.0605 | -28.036
18.0076 | -35.061 | -28.043
18.0137 | -35.0615 | -28.0503
18.0198 | -35.062 | -28.0575
18.0259 | -35.0625 | -28.0648
18.032 | -35.0629 | -28.0722
18.0381 | -35.0634 | -28.0797
18.0442 | -35.0639 | -28.0869
18.0503 | -35.0644 | -28.0943
18.0564 | -35.0649 | -28.1015
18.0625 | -35.0654 | -28.1089
18.0686 | -35.0658 | -28.1163
18.0747 | -35.0663 | -28.1237
18.0808 | -35.0668 | -28.1309
18.0869 | -35.0673 | -28.1383
18.093 | -35.0678 | -28.1454
18.0991 | -35.0683 | -28.1527
18.1052 | -35.0687 | -28.16
18.1113 | -35.0692 | -28.1672
18.1174 | -35.0697 | -28.1742
18.1236 | -35.0702 | -28.1816
18.1297 | -35.0707 | -28.189
18.1358 | -35.0712 | -28.1962
18.1419 | -35.0717 | -28.2035
18.148 | -35.0721 | -28.2108
18.1541 | -35.0726 | -28.2177
18.1602 | -35.0731 | -28.225
18.1663 | -35.0736 | -28.2322
18.1724 | -35.0741 | -28.2393
18.1785 | -35.0746 | -28.2465
18.1846 | -35.0751 | -28.2539
18.1907 | -35.0755 | -28.2612
18.1968 | -35.076 | -28.2684
18.2029 | -35.0765 | -28.2758
18.209 | -35.077 | -28.2832
18.2151 | -35.0775 | -28.2906
18.2212 | -35.078 | -28.2977
18.2273 | -35.0785 | -28.305
18.2334 | -35.079 | -28.3122
18.2395 | -35.0795 | -28.3193
18.2456 | -35.08 | -28.3267
18.2517 | -35.0804 | -28.334
18.2578 | -35.0809 | -28.341
18.2639 | -35.0814 | -28.3483
18.2701 | -35.0819 | -28.3557
18.2762 | -35.0824 | -28.3628
18.2823 | -35.0829 | -28.37
18.2884 | -35.0833 | -28.3773
18.2945 | -35.0838 | -28.3844
18.3006 | -35.0843 | -28.3918
18.3067 | -35.0848 | -28.3991
18.3128 | -35.0853 | -28.4065
18.3189 | -35.0858 | -28.4137
18.325 | -35.0863 | -28.4206
18.3311 | -35.0867 | -28.4279
18.3372 | -35.0872 | -28.4352
18.3433 | -35.0877 | -28.4426
18.3494 | -35.0882 | -28.4498
18.3555 | -35.0887 | -28.4572
18.3616 | -35.0892 | -28.4642
18.3677 | -35.0896 | -28.4717
18.3738 | -35.0901 | -28.4787
18.3799 | -35.0906 | -28.4859
18.386 | -35.0911 | -28.4931
18.3921 | -35.0915 | -28.5004
18.3982 | -35.092 | -28.5075
18.4043 | -35.0925 | -28.5148
18.4105 | -35.0929 | -28.5221
18.4166 | -35.0934 | -28.5294
18.4227 | -35.0939 | -28.5366
18.4288 | -35.0944 | -28.5439
18.4349 | -35.0949 | -28.551
18.441 | -35.0954 | -28.5582
18.4471 | -35.0958 | -28.5653
18.4532 | -35.0963 | -28.5725
18.4593 | -35.0968 | -28.5798
18.4654 | -35.0972 | -28.5867
18.4715 | -35.0977 | -28.5941
18.4776 | -35.0982 | -28.6014
18.4837 | -35.0986 | -28.6087
18.4898 | -35.0991 | -28.616
18.4959 | -35.0996 | -28.6233
18.502 | -35.1 | -28.6303
18.5081 | -35.1005 | -28.6377
18.5142 | -35.1009 | -28.645
18.5203 | -35.1014 | -28.6523
18.5264 | -35.1019 | -28.6595
18.5325 | -35.1024 | -28.6668
18.5386 | -35.1028 | -28.6739
18.5447 | -35.1033 | -28.6811
18.5508 | -35.1037 | -28.6883
18.557 | -35.1042 | -28.6957
18.5631 | -35.1047 | -28.7031
18.5692 | -35.1051 | -28.7102
18.5753 | -35.1056 | -28.7172
18.5814 | -35.1061 | -28.7244
18.5875 | -35.1066 | -28.7316
18.5936 | -35.107 | -28.7389
18.5997 | -35.1075 | -28.7461
18.6058 | -35.1079 | -28.753
18.6119 | -35.1084 | -28.7602
18.618 | -35.1089 | -28.7675
18.6241 | -35.1093 | -28.7748
18.6302 | -35.1098 | -28.782
18.6363 | -35.1103 | -28.7892
18.6424 | -35.1108 | -28.7964
18.6485 | -35.1113 | -28.8036
18.6546 | -35.1117 | -28.8109
18.6607 | -35.1122 | -28.818
18.6668 | -35.1127 | -28.8252
18.6729 | -35.1132 | -28.8324
18.679 | -35.1137 | -28.8394
18.6851 | -35.1141 | -28.8466
18.6912 | -35.1146 | -28.8536
18.6974 | -35.1151 | -28.8607
18.7035 | -35.1156 | -28.868
18.7096 | -35.116 | -28.8751
18.7157 | -35.1165 | -28.8821
18.7218 | -35.117 | -28.8895
18.7279 | -35.1174 | -28.8966
18.734 | -35.1179 | -28.9038
18.7401 | -35.1184 | -28.9109
18.7462 | -35.1189 | -28.9179
18.7523 | -35.1194 | -28.925
18.7584 | -35.1199 | -28.9324
18.7645 | -35.1203 | -28.9396
18.7706 | -35.1208 | -28.9467
18.7767 | -35.1212 | -28.9539
18.7828 | -35.1217 | -28.9608
18.7889 | -35.1222 | -28.9681
18.795 | -35.1227 | -28.9753
18.8011 | -35.1231 | -28.9825
18.8072 | -35.1236 | -28.9898
18.8133 | -35.124 | -28.9969
18.8194 | -35.1245 | -29.0039
18.8255 | -35.1249 | -29.0111
18.8316 | -35.1254 | -29.0183
18.8377 | -35.1258 | -29.0253
18.8439 | -35.1263 | -29.0325
18.85 | -35.1268 | -29.0396
18.8561 | -35.1273 | -29.0466
18.8622 | -35.1277 | -29.054
18.8683 | -35.1282 | -29.0611
18.8744 | -35.1287 | -29.0682
18.8805 | -35.1292 | -29.0754
18.8866 | -35.1297 | -29.0826
18.8927 | -35.1302 | -29.0895
18.8988 | -35.1306 | -29.0965
18.9049 | -35.1311 | -29.1037
18.911 | -35.1316 | -29.111
18.9171 | -35.1321 | -29.1182
18.9232 | -35.1325 | -29.1255
18.9293 | -35.1331 | -29.1328
18.9354 | -35.1335 | -29.1399
18.9415 | -35.134 | -29.1473
18.9476 | -35.1345 | -29.1544
18.9537 | -35.1349 | -29.1617
18.9598 | -35.1354 | -29.1687
18.9659 | -35.1358 | -29.1757
18.972 | -35.1363 | -29.1829
18.9781 | -35.1367 | -29.1901
18.9843 | -35.1372 | -29.1973
18.9904 | -35.1377 | -29.2045
18.9965 | -35.1381 | -29.2113
19.0026 | -35.1386 | -29.2185
19.0087 | -35.1391 | -29.2255
19.0148 | -35.1396 | -29.2328
19.0209 | -35.14 | -29.2401
19.027 | -35.1405 | -29.2474
19.0331 | -35.141 | -29.2541
19.0392 | -35.1415 | -29.2612
19.0453 | -35.142 | -29.2684
19.0514 | -35.1424 | -29.2752
19.0575 | -35.1429 | -29.2825
19.0636 | -35.1433 | -29.2895
19.0697 | -35.1438 | -29.2965
19.0758 | -35.1442 | -29.3036
19.0819 | -35.1447 | -29.3105
19.088 | -35.1452 | -29.3175
19.0941 | -35.1456 | -29.3247
19.1002 | -35.1461 | -29.3317
19.1063 | -35.1465 | -29.3387
19.1124 | -35.147 | -29.3457
19.1185 | -35.1474 | -29.3527
19.1246 | -35.1479 | -29.3597
19.1308 | -35.1484 | -29.3669
19.1369 | -35.1488 | -29.3739
19.143 | -35.1493 | -29.3809
19.1491 | -35.1497 | -29.3881
19.1552 | -35.1502 | -29.3953
19.1613 | -35.1507 | -29.4026
19.1674 | -35.1511 | -29.4097
19.1735 | -35.1516 | -29.4165
19.1796 | -35.152 | -29.4236
19.1857 | -35.1525 | -29.4309
19.1918 | -35.1529 | -29.4381
19.1979 | -35.1534 | -29.445
19.204 | -35.1539 | -29.4517
19.2101 | -35.1544 | -29.4586
19.2162 | -35.1549 | -29.4656
19.2223 | -35.1553 | -29.4725
19.2284 | -35.1557 | -29.4795
19.2345 | -35.1562 | -29.4865
19.2406 | -35.1566 | -29.4934
19.2467 | -35.1571 | -29.5006
19.2528 | -35.1576 | -29.5075
19.2589 | -35.158 | -29.5144
19.265 | -35.1585 | -29.5217
19.2712 | -35.159 | -29.5288
19.2773 | -35.1594 | -29.536
19.2834 | -35.1599 | -29.5431
19.2895 | -35.1604 | -29.55
19.2956 | -35.1609 | -29.5572
19.3017 | -35.1613 | -29.5641
19.3078 | -35.1618 | -29.5709
19.3139 | -35.1622 | -29.578
19.32 | -35.1627 | -29.5848
19.3261 | -35.1631 | -29.5918
19.3322 | -35.1636 | -29.5987
19.3383 | -35.1641 | -29.6057
19.3444 | -35.1646 | -29.6125
19.3505 | -35.165 | -29.6195
19.3566 | -35.1654 | -29.6264
19.3627 | -35.1659 | -29.6333
19.3688 | -35.1664 | -29.6402
19.3749 | -35.1668 | -29.6474
19.381 | -35.1673 | -29.6544
19.3871 | -35.1678 | -29.6614
19.3932 | -35.1682 | -29.6686
19.3993 | -35.1687 | -29.6756
19.4054 | -35.1692 | -29.6825
19.4115 | -35.1696 | -29.6898
19.4177 | -35.1701 | -29.6967
19.4238 | -35.1706 | -29.7037
19.4299 | -35.171 | -29.7106
19.436 | -35.1715 | -29.7178
19.4421 | -35.1719 | -29.7249
19.4482 | -35.1724 | -29.7319
19.4543 | -35.1729 | -29.7384
19.4604 | -35.1733 | -29.7454
19.4665 | -35.1738 | -29.7524
19.4726 | -35.1742 | -29.759
19.4787 | -35.1747 | -29.7658
19.4848 | -35.1751 | -29.7725
19.4909 | -35.1756 | -29.7792
19.497 | -35.176 | -29.786
19.5031 | -35.1765 | -29.793
19.5092 | -35.177 | -29.7999
19.5153 | -35.1774 | -29.8069
19.5214 | -35.1779 | -29.8137
19.5275 | -35.1783 | -29.8206
19.5336 | -35.1787 | -29.8275
19.5397 | -35.1792 | -29.8344
19.5458 | -35.1796 | -29.8414
19.5519 | -35.1801 | -29.8483
19.5581 | -35.1805 | -29.8551
19.5642 | -35.181 | -29.8621
19.5703 | -35.1814 | -29.869
19.5764 | -35.1819 | -29.8759
19.5825 | -35.1823 | -29.883
19.5886 | -35.1828 | -29.89
19.5947 | -35.1832 | -29.8967
19.6008 | -35.1836 | -29.9035
19.6069 | -35.1841 | -29.9102
19.613 | -35.1845 | -29.9173
19.6191 | -35.185 | -29.9242
19.6252 | -35.1855 | -29.9311
19.6313 | -35.1859 | -29.9377
19.6374 | -35.1864 | -29.9446
19.6435 | -35.1868 | -29.9514
19.6496 | -35.1873 | -29.9581
19.6557 | -35.1878 | -29.965
19.6618 | -35.1882 | -29.972
19.6679 | -35.1887 | -29.9791
19.674 | -35.1892 | -29.9859
19.6801 | -35.1896 | -29.9929
19.6862 | -35.1901 | -29.9998
19.6923 | -35.1905 | -30.0067
19.6984 | -35.191 | -30.0136
19.7046 | -35.1914 | -30.0204
19.7107 | -35.1919 | -30.0274
19.7168 | -35.1923 | -30.0347
19.7229 | -35.1928 | -30.0415
19.729 | -35.1932 | -30.0487
19.7351 | -35.1936 | -30.0552
19.7412 | -35.1941 | -30.0623
19.7473 | -35.1945 | -30.0693
19.7534 | -35.1949 | -30.0763
19.7595 | -35.1954 | -30.0831
19.7656 | -35.1958 | -30.0902
19.7717 | -35.1963 | -30.0971
19.7778 | -35.1967 | -30.104
19.7839 | -35.1972 | -30.1108
19.79 | -35.1976 | -30.1176
19.7961 | -35.198 | -30.1242
19.8022 | -35.1985 | -30.1308
19.8083 | -35.199 | -30.1377
19.8144 | -35.1994 | -30.1446
19.8205 | -35.1999 | -30.1515
19.8266 | -35.2004 | -30.1583
19.8327 | -35.2008 | -30.1652
19.8388 | -35.2012 | -30.1718
19.845 | -35.2017 | -30.1786
19.8511 | -35.2021 | -30.1853
19.8572 | -35.2025 | -30.1921
19.8633 | -35.203 | -30.1992
19.8694 | -35.2035 | -30.2065
19.8755 | -35.2039 | -30.2131
19.8816 | -35.2044 | -30.2203
19.8877 | -35.2048 | -30.227
19.8938 | -35.2053 | -30.234
19.8999 | -35.2057 | -30.2407
19.906 | -35.2062 | -30.2474
19.9121 | -35.2066 | -30.2544
19.9182 | -35.207 | -30.2615
19.9243 | -35.2075 | -30.2683
19.9304 | -35.2079 | -30.2752
19.9365 | -35.2084 | -30.282
19.9426 | -35.2089 | -30.2882
19.9487 | -35.2094 | -30.2951
19.9548 | -35.2098 | -30.3021
19.9609 | -35.2103 | -30.309
19.967 | -35.2108 | -30.3159
19.9731 | -35.2112 | -30.3231
19.9792 | -35.2117 | -30.3296
19.9853 | -35.2121 | -30.3361
19.9915 | -35.2126 | -30.343
19.9976 | -35.2131 | -30.3498
20.0037 | -35.2136 | -30.3568
20.0098 | -35.214 | -30.364
20.0159 | -35.2145 | -30.3706
20.022 | -35.215 | -30.3776
20.0281 | -35.2154 | -30.3845
20.0342 | -35.2159 | -30.3915
20.0403 | -35.2163 | -30.3988
20.0464 | -35.2168 | -30.4053
20.0525 | -35.2172 | -30.4124
20.0586 | -35.2177 | -30.4191
20.0647 | -35.2181 | -30.4261
20.0708 | -35.2186 | -30.4332
20.0769 | -35.219 | -30.4398
20.083 | -35.2195 | -30.4465
20.0891 | -35.22 | -30.453
20.0952 | -35.2205 | -30.4598
20.1013 | -35.2209 | -30.4663
20.1074 | -35.2213 | -30.4732
20.1135 | -35.2218 | -30.4802
20.1196 | -35.2222 | -30.4872
20.1257 | -35.2226 | -30.4939
20.1319 | -35.223 | -30.5009
20.138 | -35.2235 | -30.5073
20.1441 | -35.2239 | -30.5142
20.1502 | -35.2244 | -30.5208
20.1563 | -35.2249 | -30.5273
20.1624 | -35.2253 | -30.5341
20.1685 | -35.2258 | -30.5414
20.1746 | -35.2263 | -30.5482
20.1807 | -35.2267 | -30.5547
20.1868 | -35.2271 | -30.5615
20.1929 | -35.2275 | -30.5686
20.199 | -35.228 | -30.5754
20.2051 | -35.2285 | -30.5822
20.2112 | -35.229 | -30.5893
20.2173 | -35.2294 | -30.5959
20.2234 | -35.2299 | -30.6026
20.2295 | -35.2304 | -30.6092
20.2356 | -35.2308 | -30.616
20.2417 | -35.2313 | -30.6225
20.2478 | -35.2317 | -30.629
20.2539 | -35.2321 | -30.6356
20.26 | -35.2326 | -30.6428
20.2661 | -35.2331 | -30.6495
20.2723 | -35.2335 | -30.6564
20.2784 | -35.234 | -30.663
20.2845 | -35.2344 | -30.6699
20.2906 | -35.2349 | -30.6766
20.2967 | -35.2353 | -30.6832
20.3028 | -35.2357 | -30.6901
20.3089 | -35.2362 | -30.6966
20.315 | -35.2366 | -30.7033
20.3211 | -35.2371 | -30.7099
20.3272 | -35.2376 | -30.7167
20.3333 | -35.238 | -30.7233
20.3394 | -35.2385 | -30.7297
20.3455 | -35.2388 | -30.7368
20.3516 | -35.2392 | -30.7433
20.3577 | -35.2397 | -30.7499
20.3638 | -35.2401 | -30.7566
20.3699 | -35.2406 | -30.7629
20.376 | -35.2411 | -30.7695
20.3821 | -35.2415 | -30.7762
20.3882 | -35.242 | -30.7829
20.3943 | -35.2424 | -30.7893
20.4004 | -35.2429 | -30.7954
20.4065 | -35.2433 | -30.8021
20.4126 | -35.2437 | -30.8089
20.4188 | -35.2442 | -30.8156
20.4249 | -35.2447 | -30.8217
20.431 | -35.2451 | -30.8286
20.4371 | -35.2456 | -30.8355
20.4432 | -35.246 | -30.8423
20.4493 | -35.2465 | -30.8489
20.4554 | -35.2469 | -30.8555
20.4615 | -35.2473 | -30.8621
20.4676 | -35.2477 | -30.8691
20.4737 | -35.2482 | -30.8757
20.4798 | -35.2487 | -30.8822
20.4859 | -35.2491 | -30.8889
20.492 | -35.2496 | -30.8956
20.4981 | -35.2501 | -30.9022
20.5042 | -35.2505 | -30.9091
20.5103 | -35.251 | -30.9153
20.5164 | -35.2514 | -30.9217
20.5225 | -35.2519 | -30.9283
20.5286 | -35.2523 | -30.9352
20.5347 | -35.2528 | -30.9419
20.5408 | -35.2532 | -30.9483
20.5469 | -35.2537 | -30.9549
20.553 | -35.2541 | -30.9618
20.5592 | -35.2546 | -30.9685
20.5653 | -35.255 | -30.9752
20.5714 | -35.2555 | -30.9816
20.5775 | -35.2559 | -30.9885
20.5836 | -35.2563 | -30.9954
20.5897 | -35.2568 | -31.0024
20.5958 | -35.2572 | -31.0092
20.6019 | -35.2577 | -31.0161
20.608 | -35.2581 | -31.023
20.6141 | -35.2586 | -31.0296
20.6202 | -35.2591 | -31.0361
20.6263 | -35.2595 | -31.0426
20.6324 | -35.26 | -31.0493
20.6385 | -35.2605 | -31.0559
20.6446 | -35.2609 | -31.0624
20.6507 | -35.2614 | -31.0686
20.6568 | -35.2619 | -31.0751
20.6629 | -35.2623 | -31.0819
20.669 | -35.2628 | -31.0887
20.6751 | -35.2632 | -31.095
20.6812 | -35.2636 | -31.1013
20.6873 | -35.264 | -31.1079
20.6934 | -35.2645 | -31.1145
20.6995 | -35.265 | -31.121
20.7057 | -35.2654 | -31.1278
20.7118 | -35.2659 | -31.1343
20.7179 | -35.2663 | -31.1406
20.724 | -35.2668 | -31.1475
20.7301 | -35.2672 | -31.1539
20.7362 | -35.2677 | -31.1606
20.7423 | -35.2682 | -31.1673
20.7484 | -35.2687 | -31.1742
20.7545 | -35.2692 | -31.1809
20.7606 | -35.2696 | -31.1876
20.7667 | -35.2701 | -31.1943
20.7728 | -35.2706 | -31.201
20.7789 | -35.271 | -31.2075
20.785 | -35.2715 | -31.2142
20.7911 | -35.272 | -31.221
20.7972 | -35.2724 | -31.2274
20.8033 | -35.2729 | -31.2342
20.8094 | -35.2733 | -31.2406
20.8155 | -35.2737 | -31.2472
20.8216 | -35.2741 | -31.2536
20.8277 | -35.2746 | -31.2602
20.8338 | -35.2751 | -31.2671
20.8399 | -35.2755 | -31.2737
20.8461 | -35.276 | -31.2805
20.8522 | -35.2764 | -31.287
20.8583 | -35.2769 | -31.2937
20.8644 | -35.2773 | -31.3006
20.8705 | -35.2778 | -31.3072
20.8766 | -35.2782 | -31.3139
20.8827 | -35.2787 | -31.3201
20.8888 | -35.2791 | -31.3267
20.8949 | -35.2796 | -31.333
20.901 | -35.28 | -31.3395
20.9071 | -35.2804 | -31.3459
20.9132 | -35.2809 | -31.352
20.9193 | -35.2813 | -31.3584
20.9254 | -35.2818 | -31.365
20.9315 | -35.2823 | -31.3715
20.9376 | -35.2827 | -31.378
20.9437 | -35.2831 | -31.3844
20.9498 | -35.2836 | -31.3908
20.9559 | -35.2841 | -31.3974
20.962 | -35.2846 | -31.4038
20.9681 | -35.285 | -31.41
20.9742 | -35.2855 | -31.4164
20.9803 | -35.2859 | -31.4229
20.9864 | -35.2863 | -31.4296
20.9926 | -35.2868 | -31.4365
20.9987 | -35.2872 | -31.4425
21.0048 | -35.2877 | -31.4492
21.0109 | -35.2882 | -31.4558
21.017 | -35.2886 | -31.4622
21.0231 | -35.2891 | -31.469
21.0292 | -35.2895 | -31.4756
21.0353 | -35.29 | -31.4822
21.0414 | -35.2904 | -31.4886
21.0475 | -35.2909 | -31.4949
21.0536 | -35.2914 | -31.5015
21.0597 | -35.2918 | -31.5081
21.0658 | -35.2922 | -31.5148
21.0719 | -35.2927 | -31.5212
21.078 | -35.2931 | -31.5282
21.0841 | -35.2936 | -31.5349
21.0902 | -35.294 | -31.5412
21.0963 | -35.2944 | -31.5482
21.1024 | -35.2948 | -31.5543
21.1085 | -35.2952 | -31.5606
21.1146 | -35.2957 | -31.567
21.1207 | -35.2962 | -31.5733
21.1268 | -35.2966 | -31.5801
21.133 | -35.2971 | -31.5865
21.1391 | -35.2975 | -31.5926
21.1452 | -35.298 | -31.599
21.1513 | -35.2984 | -31.6057
21.1574 | -35.2989 | -31.6121
21.1635 | -35.2994 | -31.6188
21.1696 | -35.2998 | -31.6251
21.1757 | -35.3002 | -31.6313
21.1818 | -35.3007 | -31.6379
21.1879 | -35.3011 | -31.6445
21.194 | -35.3015 | -31.6507
21.2001 | -35.302 | -31.6575
21.2062 | -35.3024 | -31.664
21.2123 | -35.3029 | -31.6705
21.2184 | -35.3033 | -31.6777
21.2245 | -35.3038 | -31.6837
21.2306 | -35.3042 | -31.6899
21.2367 | -35.3047 | -31.6964
21.2428 | -35.3051 | -31.7027
21.2489 | -35.3055 | -31.7093
21.255 | -35.306 | -31.7158
21.2611 | -35.3064 | -31.7222
21.2672 | -35.3069 | -31.7286
21.2733 | -35.3073 | -31.7351
21.2795 | -35.3078 | -31.7413
21.2856 | -35.3082 | -31.7477
21.2917 | -35.3087 | -31.7541
21.2978 | -35.3091 | -31.7604
21.3039 | -35.3096 | -31.7669
21.31 | -35.31 | -31.7736
21.3161 | -35.3105 | -31.7803
21.3222 | -35.311 | -31.7867
21.3283 | -35.3114 | -31.7933
21.3344 | -35.3118 | -31.7995
21.3405 | -35.3122 | -31.8062
21.3466 | -35.3127 | -31.8125
21.3527 | -35.3131 | -31.8184
21.3588 | -35.3136 | -31.8249
21.3649 | -35.314 | -31.8312
21.371 | -35.3145 | -31.838
21.3771 | -35.3149 | -31.8441
21.3832 | -35.3154 | -31.8507
21.3893 | -35.3159 | -31.8568
21.3954 | -35.3163 | -31.8633
21.4015 | -35.3168 | -31.8697
21.4076 | -35.3172 | -31.8757
21.4137 | -35.3177 | -31.882
21.4199 | -35.3182 | -31.8884
21.426 | -35.3186 | -31.8945
21.4321 | -35.3191 | -31.9011
21.4382 | -35.3195 | -31.9076
21.4443 | -35.3199 | -31.914
21.4504 | -35.3204 | -31.9204
21.4565 | -35.3209 | -31.9268
21.4626 | -35.3213 | -31.9333
21.4687 | -35.3218 | -31.9399
21.4748 | -35.3222 | -31.9462
21.4809 | -35.3227 | -31.9526
21.487 | -35.3231 | -31.959
21.4931 | -35.3235 | -31.9651
21.4992 | -35.3239 | -31.9715
21.5053 | -35.3243 | -31.9778
21.5114 | -35.3248 | -31.9845
21.5175 | -35.3252 | -31.9911
21.5236 | -35.3256 | -31.9975
21.5297 | -35.3261 | -32.0031
21.5358 | -35.3265 | -32.0093
21.5419 | -35.327 | -32.0156
21.548 | -35.3274 | -32.0219
21.5541 | -35.3279 | -32.0283
21.5602 | -35.3284 | -32.0347
21.5664 | -35.3288 | -32.0408
21.5725 | -35.3292 | -32.0471
21.5786 | -35.3296 | -32.0535
21.5847 | -35.33 | -32.0597
21.5908 | -35.3305 | -32.066
21.5969 | -35.3309 | -32.0722
21.603 | -35.3314 | -32.0786
21.6091 | -35.3318 | -32.0849
21.6152 | -35.3323 | -32.0916
21.6213 | -35.3327 | -32.098
21.6274 | -35.3332 | -32.1046
21.6335 | -35.3336 | -32.1107
21.6396 | -35.334 | -32.1172
21.6457 | -35.3344 | -32.1235
21.6518 | -35.3348 | -32.1298
21.6579 | -35.3353 | -32.1361
21.664 | -35.3358 | -32.1424
21.6701 | -35.3361 | -32.1483
21.6762 | -35.3366 | -32.1544
21.6823 | -35.337 | -32.1605
21.6884 | -35.3375 | -32.1671
21.6945 | -35.3379 | -32.1734
21.7006 | -35.3384 | -32.1796
21.7068 | -35.3389 | -32.1855
21.7129 | -35.3394 | -32.1917
21.719 | -35.3399 | -32.1985
21.7251 | -35.3403 | -32.205
21.7312 | -35.3408 | -32.2112
21.7373 | -35.3412 | -32.2168
21.7434 | -35.3417 | -32.2234
21.7495 | -35.3422 | -32.2296
21.7556 | -35.3426 | -32.2357
21.7617 | -35.3431 | -32.2418
21.7678 | -35.3435 | -32.2477
21.7739 | -35.344 | -32.253
21.78 | -35.3445 | -32.2594
21.7861 | -35.3449 | -32.2657
21.7922 | -35.3453 | -32.272
21.7983 | -35.3457 | -32.2782
21.8044 | -35.3462 | -32.2847
21.8105 | -35.3467 | -32.2911
21.8166 | -35.3471 | -32.2976
21.8227 | -35.3475 | -32.3037
21.8288 | -35.348 | -32.3099
21.8349 | -35.3484 | -32.3162
21.841 | -35.3489 | -32.3224
21.8471 | -35.3493 | -32.3283
21.8533 | -35.3498 | -32.3343
21.8594 | -35.3501 | -32.3407
21.8655 | -35.3506 | -32.3465
21.8716 | -35.351 | -32.3532
21.8777 | -35.3515 | -32.3597
21.8838 | -35.3519 | -32.3659
21.8899 | -35.3523 | -32.3719
21.896 | -35.3528 | -32.3786
21.9021 | -35.3533 | -32.3853
21.9082 | -35.3538 | -32.3918
21.9143 | -35.3542 | -32.3979
21.9204 | -35.3547 | -32.4037
21.9265 | -35.3552 | -32.4099
21.9326 | -35.3556 | -32.416
21.9387 | -35.356 | -32.4226
21.9448 | -35.3564 | -32.429
21.9509 | -35.3568 | -32.4347
21.957 | -35.3573 | -32.4411
21.9631 | -35.3579 | -32.4473
21.9692 | -35.3583 | -32.4538
21.9753 | -35.3587 | -32.4598
21.9814 | -35.3592 | -32.4658
21.9875 | -35.3596 | -32.4719
21.9937 | -35.36 | -32.4782
21.9998 | -35.3605 | -32.4841
22.0059 | -35.3609 | -32.4909
22.012 | -35.3614 | -32.4976
22.0181 | -35.3617 | -32.5037
22.0242 | -35.3622 | -32.5098
22.0303 | -35.3626 | -32.5158
22.0364 | -35.3631 | -32.522
22.0425 | -35.3636 | -32.5284
22.0486 | -35.364 | -32.5345
22.0547 | -35.3645 | -32.5402
22.0608 | -35.365 | -32.5467
22.0669 | -35.3655 | -32.5533
22.073 | -35.3659 | -32.5593
22.0791 | -35.3663 | -32.5656
22.0852 | -35.3668 | -32.5719
22.0913 | -35.3672 | -32.5785
22.0974 | -35.3677 | -32.585
22.1035 | -35.3681 | -32.591
22.1096 | -35.3686 | -32.5974
22.1157 | -35.3691 | -32.6039
22.1218 | -35.3695 | -32.6099
22.1279 | -35.37 | -32.6156
22.134 | -35.3705 | -32.6216
22.1402 | -35.371 | -32.6271
22.1463 | -35.3714 | -32.6331
22.1524 | -35.372 | -32.6392
22.1585 | -35.3723 | -32.6458
22.1646 | -35.3728 | -32.6517
22.1707 | -35.3732 | -32.6578
22.1768 | -35.3737 | -32.6642
22.1829 | -35.3741 | -32.6701
22.189 | -35.3746 | -32.6761
22.1951 | -35.375 | -32.6819
22.2012 | -35.3754 | -32.6882
22.2073 | -35.3759 | -32.6946
22.2134 | -35.3763 | -32.7011
22.2195 | -35.3767 | -32.7076
22.2256 | -35.3772 | -32.7133
22.2317 | -35.3778 | -32.7199
22.2378 | -35.3782 | -32.7262
22.2439 | -35.3788 | -32.7329
22.25 | -35.3791 | -32.7392
22.2561 | -35.3796 | -32.7448
22.2622 | -35.38 | -32.751
22.2683 | -35.3804 | -32.7569
22.2744 | -35.3808 | -32.7626
22.2806 | -35.3812 | -32.7688
22.2867 | -35.3817 | -32.7751
22.2928 | -35.3821 | -32.7812
22.2989 | -35.3826 | -32.7873
22.305 | -35.383 | -32.7932
22.3111 | -35.3834 | -32.7995
22.3172 | -35.3839 | -32.8055
22.3233 | -35.3843 | -32.8112
22.3294 | -35.3847 | -32.8175
22.3355 | -35.3853 | -32.8236
22.3416 | -35.3857 | -32.8296
22.3477 | -35.3862 | -32.8363
22.3538 | -35.3866 | -32.8424
22.3599 | -35.3871 | -32.8487
22.366 | -35.3876 | -32.8548
22.3721 | -35.388 | -32.8609
22.3782 | -35.3884 | -32.867
22.3843 | -35.3888 | -32.873
22.3904 | -35.3893 | -32.8789
22.3965 | -35.3897 | -32.8853
22.4026 | -35.3902 | -32.8913
22.4087 | -35.3907 | -32.8971
22.4148 | -35.3911 | -32.9036
22.4209 | -35.3916 | -32.9094
22.4271 | -35.392 | -32.9154
22.4332 | -35.3924 | -32.9215
22.4393 | -35.3929 | -32.9276
22.4454 | -35.3933 | -32.9332
22.4515 | -35.3938 | -32.9397
22.4576 | -35.3942 | -32.9462
22.4637 | -35.3946 | -32.9523
22.4698 | -35.3951 | -32.9586
22.4759 | -35.3956 | -32.9646
22.482 | -35.396 | -32.9706
22.4881 | -35.3965 | -32.9767
22.4942 | -35.3969 | -32.9831
22.5003 | -35.3973 | -32.9896
22.5064 | -35.3978 | -32.9956
22.5125 | -35.3982 | -33.0012
22.5186 | -35.3986 | -33.0075
22.5247 | -35.399 | -33.0134
22.5308 | -35.3995 | -33.0195
22.5369 | -35.3999 | -33.0255
22.543 | -35.4003 | -33.0315
22.5491 | -35.4008 | -33.0371
22.5552 | -35.4012 | -33.0434
22.5613 | -35.4017 | -33.0498
22.5675 | -35.4021 | -33.0563
22.5736 | -35.4025 | -33.0623
22.5797 | -35.4029 | -33.0681
22.5858 | -35.4033 | -33.0741
22.5919 | -35.4038 | -33.0802
22.598 | -35.4042 | -33.0862
22.6041 | -35.4046 | -33.0923
22.6102 | -35.405 | -33.0983
22.6163 | -35.4055 | -33.1039
22.6224 | -35.4059 | -33.1099
22.6285 | -35.4063 | -33.1159
22.6346 | -35.4068 | -33.122
22.6407 | -35.4073 | -33.128
22.6468 | -35.4077 | -33.1337
22.6529 | -35.4081 | -33.1393
22.659 | -35.4086 | -33.1454
22.6651 | -35.409 | -33.1517
22.6712 | -35.4094 | -33.1577
22.6773 | -35.4099 | -33.1635
22.6834 | -35.4103 | -33.1699
22.6895 | -35.4107 | -33.176
22.6956 | -35.4112 | -33.1823
22.7017 | -35.4116 | -33.1881
22.7079 | -35.412 | -33.1941
22.714 | -35.4125 | -33.2005
22.7201 | -35.4129 | -33.2068
22.7262 | -35.4133 | -33.2122
22.7323 | -35.4138 | -33.2181
22.7384 | -35.4142 | -33.2245
22.7445 | -35.4146 | -33.2302
22.7506 | -35.415 | -33.2362
22.7567 | -35.4154 | -33.2422
22.7628 | -35.4159 | -33.2484
22.7689 | -35.4163 | -33.2542
22.775 | -35.4168 | -33.2608
22.7811 | -35.4172 | -33.2661
22.7872 | -35.4176 | -33.2723
22.7933 | -35.4181 | -33.2785
22.7994 | -35.4185 | -33.2843
22.8055 | -35.419 | -33.2903
22.8116 | -35.4194 | -33.2961
22.8177 | -35.4199 | -33.3023
22.8238 | -35.4203 | -33.3081
22.8299 | -35.4207 | -33.3136
22.836 | -35.4211 | -33.3197
22.8421 | -35.4216 | -33.3262
22.8482 | -35.4219 | -33.3322
22.8544 | -35.4224 | -33.3384
22.8605 | -35.4229 | -33.3446
22.8666 | -35.4232 | -33.3501
22.8727 | -35.4237 | -33.3564
22.8788 | -35.4242 | -33.3626
22.8849 | -35.4246 | -33.3689
22.891 | -35.4251 | -33.3745
22.8971 | -35.4255 | -33.3802
22.9032 | -35.4259 | -33.386
22.9093 | -35.4263 | -33.392
22.9154 | -35.4268 | -33.3982
22.9215 | -35.4272 | -33.4042
22.9276 | -35.4277 | -33.4102
22.9337 | -35.4281 | -33.4164
22.9398 | -35.4286 | -33.4224
22.9459 | -35.4291 | -33.4289
22.952 | -35.4295 | -33.4343
22.9581 | -35.4299 | -33.4407
22.9642 | -35.4303 | -33.447
22.9703 | -35.4308 | -33.4526
22.9764 | -35.4312 | -33.4584
22.9825 | -35.4316 | -33.4646
22.9886 | -35.432 | -33.4709
22.9948 | -35.4324 | -33.4774
23.0009 | -35.4328 | -33.4836
23.007 | -35.4333 | -33.4898
23.0131 | -35.4337 | -33.4959
23.0192 | -35.4343 | -33.5022
23.0253 | -35.4347 | -33.508
23.0314 | -35.4352 | -33.5144
23.0375 | -35.4356 | -33.5199
23.0436 | -35.436 | -33.5254
23.0497 | -35.4364 | -33.5316
23.0558 | -35.4369 | -33.5371
23.0619 | -35.4373 | -33.5434
23.068 | -35.4378 | -33.5491
23.0741 | -35.4382 | -33.5554
23.0802 | -35.4386 | -33.5614
23.0863 | -35.4391 | -33.5676
23.0924 | -35.4395 | -33.5734
23.0985 | -35.44 | -33.5794
23.1046 | -35.4404 | -33.5852
23.1107 | -35.4408 | -33.5907
23.1168 | -35.4412 | -33.5963
23.1229 | -35.4416 | -33.6022
23.129 | -35.442 | -33.6081
23.1351 | -35.4424 | -33.6141
23.1413 | -35.4428 | -33.6203
23.1474 | -35.4432 | -33.6252
23.1535 | -35.4437 | -33.6316
23.1596 | -35.4441 | -33.6378
23.1657 | -35.4446 | -33.6441
23.1718 | -35.4449 | -33.6503
23.1779 | -35.4454 | -33.6563
23.184 | -35.4458 | -33.6623
23.1901 | -35.4463 | -33.6682
23.1962 | -35.4467 | -33.6742
23.2023 | -35.4471 | -33.68
23.2084 | -35.4476 | -33.6862
23.2145 | -35.4479 | -33.6923
23.2206 | -35.4483 | -33.698
23.2267 | -35.4487 | -33.704
23.2328 | -35.4491 | -33.7093
23.2389 | -35.4494 | -33.7153
23.245 | -35.4499 | -33.7208
23.2511 | -35.4504 | -33.7268
23.2572 | -35.4508 | -33.7327
23.2633 | -35.4513 | -33.7387
23.2694 | -35.4517 | -33.7448
23.2755 | -35.4521 | -33.7507
23.2817 | -35.4526 | -33.7567
23.2878 | -35.4529 | -33.7623
23.2939 | -35.4534 | -33.7684
23.3 | -35.4538 | -33.7739
23.3061 | -35.4542 | -33.78
23.3122 | -35.4546 | -33.7861
23.3183 | -35.455 | -33.7919
23.3244 | -35.4555 | -33.7973
23.3305 | -35.4558 | -33.8035
23.3366 | -35.4562 | -33.8098
23.3427 | -35.4566 | -33.8157
23.3488 | -35.4571 | -33.8216
23.3549 | -35.4575 | -33.828
23.361 | -35.458 | -33.8342
23.3671 | -35.4584 | -33.8404
23.3732 | -35.4588 | -33.846
23.3793 | -35.4592 | -33.8515
23.3854 | -35.4596 | -33.8577
23.3915 | -35.4601 | -33.8637
23.3976 | -35.4605 | -33.87
23.4037 | -35.461 | -33.8756
23.4098 | -35.4614 | -33.882
23.4159 | -35.4617 | -33.8882
23.422 | -35.4622 | -33.8946
23.4282 | -35.4625 | -33.9006
23.4343 | -35.4629 | -33.9066
23.4404 | -35.4633 | -33.9126
23.4465 | -35.4637 | -33.9188
23.4526 | -35.4642 | -33.9244
23.4587 | -35.4646 | -33.9302
23.4648 | -35.465 | -33.9361
23.4709 | -35.4654 | -33.9421
23.477 | -35.4658 | -33.9482
23.4831 | -35.4663 | -33.9543
23.4892 | -35.4667 | -33.96
23.4953 | -35.4671 | -33.9661
23.5014 | -35.4674 | -33.9719
23.5075 | -35.4679 | -33.9779
23.5136 | -35.4683 | -33.9837
23.5197 | -35.4688 | -33.9899
23.5258 | -35.4692 | -33.9958
23.5319 | -35.4697 | -34.0021
23.538 | -35.4701 | -34.0077
23.5441 | -35.4706 | -34.0135
23.5502 | -35.4711 | -34.019
23.5563 | -35.4714 | -34.0253
23.5624 | -35.4719 | -34.0308
23.5686 | -35.4722 | -34.0367
23.5747 | -35.4726 | -34.0427
23.5808 | -35.473 | -34.0487
23.5869 | -35.4734 | -34.0544
23.593 | -35.4738 | -34.0604
23.5991 | -35.4743 | -34.0663
23.6052 | -35.4747 | -34.0722
23.6113 | -35.4751 | -34.0784
23.6174 | -35.4755 | -34.0848
23.6235 | -35.4759 | -34.0907
23.6296 | -35.4764 | -34.0963
23.6357 | -35.4768 | -34.1022
23.6418 | -35.4772 | -34.1074
23.6479 | -35.4776 | -34.1134
23.654 | -35.478 | -34.1193
23.6601 | -35.4784 | -34.1252
23.6662 | -35.4788 | -34.131
23.6723 | -35.4792 | -34.1368
23.6784 | -35.4796 | -34.143
23.6845 | -35.48 | -34.1486
23.6906 | -35.4804 | -34.1546
23.6967 | -35.4808 | -34.1603
23.7028 | -35.4812 | -34.1665
23.7089 | -35.4814 | -34.1726
23.7151 | -35.4818 | -34.1784
23.7212 | -35.4821 | -34.1841
23.7273 | -35.4825 | -34.19
23.7334 | -35.4829 | -34.1962
23.7395 | -35.4833 | -34.2021
23.7456 | -35.4837 | -34.208
23.7517 | -35.4841 | -34.2141
23.7578 | -35.4845 | -34.2199
23.7639 | -35.4849 | -34.2259
23.77 | -35.4854 | -34.232
23.7761 | -35.4858 | -34.2379
23.7822 | -35.4862 | -34.2431
23.7883 | -35.4866 | -34.2491
23.7944 | -35.487 | -34.2547
23.8005 | -35.4874 | -34.2606
23.8066 | -35.4878 | -34.2663
23.8127 | -35.4882 | -34.2725
23.8188 | -35.4886 | -34.2787
23.8249 | -35.4891 | -34.2847
23.831 | -35.4894 | -34.2904
23.8371 | -35.4898 | -34.2961
23.8432 | -35.4902 | -34.302
23.8493 | -35.4905 | -34.3075
23.8555 | -35.4909 | -34.3134
23.8616 | -35.4913 | -34.3188
23.8677 | -35.4918 | -34.3248
23.8738 | -35.4921 | -34.3303
23.8799 | -35.4925 | -34.3363
23.886 | -35.4929 | -34.3418
23.8921 | -35.4933 | -34.3475
23.8982 | -35.4938 | -34.3534
23.9043 | -35.4942 | -34.3591
23.9104 | -35.4946 | -34.3651
23.9165 | -35.4949 | -34.3709
23.9226 | -35.4953 | -34.3767
23.9287 | -35.4957 | -34.3827
23.9348 | -35.4961 | -34.3881
23.9409 | -35.4965 | -34.3939
23.947 | -35.4969 | -34.3998
23.9531 | -35.4973 | -34.4056
23.9592 | -35.4977 | -34.4117
23.9653 | -35.4981 | -34.4173
23.9714 | -35.4985 | -34.4229
23.9775 | -35.4989 | -34.4288
23.9836 | -35.4993 | -34.4349
23.9897 | -35.4998 | -34.4408
23.9958 | -35.5002 | -34.4467
24.002 | -35.5005 | -34.4526
24.0081 | -35.5009 | -34.4585
24.0142 | -35.5013 | -34.4641
24.0203 | -35.5016 | -34.4698
24.0264 | -35.502 | -34.4756
24.0325 | -35.5024 | -34.4813
24.0386 | -35.5028 | -34.4872
24.0447 | -35.5032 | -34.4929
24.0508 | -35.5036 | -34.499
24.0569 | -35.5039 | -34.5048
24.063 | -35.5043 | -34.5107
24.0691 | -35.5047 | -34.5163
24.0752 | -35.5051 | -34.5222
24.0813 | -35.5055 | -34.5277
24.0874 | -35.5059 | -34.5335
24.0935 | -35.5063 | -34.5393
24.0996 | -35.5066 | -34.5452
24.1057 | -35.507 | -34.5509
24.1118 | -35.5075 | -34.5568
24.1179 | -35.5079 | -34.5628
24.124 | -35.5082 | -34.5686
24.1301 | -35.5086 | -34.5744
24.1362 | -35.5091 | -34.5803
24.1424 | -35.5095 | -34.5863
24.1485 | -35.5099 | -34.592
24.1546 | -35.5102 | -34.5978
24.1607 | -35.5106 | -34.6035
24.1668 | -35.511 | -34.6091
24.1729 | -35.5114 | -34.6147
24.179 | -35.5118 | -34.6205
24.1851 | -35.5122 | -34.6264
24.1912 | -35.5125 | -34.6322
24.1973 | -35.5129 | -34.6381
24.2034 | -35.5133 | -34.6437
24.2095 | -35.5137 | -34.6497
24.2156 | -35.5141 | -34.6557
24.2217 | -35.5145 | -34.6614
24.2278 | -35.5149 | -34.6668
24.2339 | -35.5152 | -34.6726
24.24 | -35.5156 | -34.6779
24.2461 | -35.516 | -34.6836
24.2522 | -35.5164 | -34.6894
24.2583 | -35.5168 | -34.6955
24.2644 | -35.5173 | -34.7014
24.2705 | -35.5177 | -34.7074
24.2766 | -35.5181 | -34.7132
24.2827 | -35.5185 | -34.7189
24.2889 | -35.5189 | -34.7248
24.295 | -35.5193 | -34.7305
24.3011 | -35.5197 | -34.736
24.3072 | -35.5201 | -34.7414
24.3133 | -35.5205 | -34.7471
24.3194 | -35.5208 | -34.7529
24.3255 | -35.5212 | -34.7588
24.3316 | -35.5216 | -34.7645
24.3377 | -35.5221 | -34.7699
24.3438 | -35.5225 | -34.7752
24.3499 | -35.5229 | -34.7812
24.356 | -35.5233 | -34.7868
24.3621 | -35.5237 | -34.7926
24.3682 | -35.5241 | -34.7982
24.3743 | -35.5245 | -34.8041
24.3804 | -35.5249 | -34.8096
24.3865 | -35.5252 | -34.8151
24.3926 | -35.5256 | -34.8212
24.3987 | -35.526 | -34.8268
24.4048 | -35.5264 | -34.8329
24.4109 | -35.5267 | -34.8392
24.417 | -35.5271 | -34.845
24.4231 | -35.5275 | -34.8507
24.4293 | -35.5279 | -34.8566
24.4354 | -35.5283 | -34.8625
24.4415 | -35.5288 | -34.8685
24.4476 | -35.5291 | -34.8743
24.4537 | -35.5295 | -34.8799
24.4598 | -35.5299 | -34.8855
24.4659 | -35.5303 | -34.8914
24.472 | -35.5307 | -34.8973
24.4781 | -35.5311 | -34.9031
24.4842 | -35.5314 | -34.9083
24.4903 | -35.5318 | -34.9139
24.4964 | -35.5322 | -34.9195
24.5025 | -35.5326 | -34.9255
24.5086 | -35.533 | -34.931
24.5147 | -35.5334 | -34.9368
24.5208 | -35.5339 | -34.9424
24.5269 | -35.5342 | -34.9479
24.533 | -35.5347 | -34.9537
24.5391 | -35.5351 | -34.9593
24.5452 | -35.5355 | -34.965
24.5513 | -35.5359 | -34.9707
24.5574 | -35.5363 | -34.9762
24.5635 | -35.5367 | -34.9819
24.5696 | -35.5371 | -34.9876
24.5758 | -35.5375 | -34.9934
24.5819 | -35.5379 | -34.9992
24.588 | -35.5383 | -35.0046
24.5941 | -35.5386 | -35.0105
24.6002 | -35.5391 | -35.0163
24.6063 | -35.5395 | -35.0219
24.6124 | -35.5399 | -35.0277
24.6185 | -35.5403 | -35.0333
24.6246 | -35.5406 | -35.0388
24.6307 | -35.541 | -35.0443
24.6368 | -35.5414 | -35.0498
24.6429 | -35.5418 | -35.0557
24.649 | -35.5422 | -35.0614
24.6551 | -35.5425 | -35.0674
24.6612 | -35.5429 | -35.0727
24.6673 | -35.5433 | -35.0785
24.6734 | -35.5437 | -35.0841
24.6795 | -35.5442 | -35.0899
24.6856 | -35.5446 | -35.0957
24.6917 | -35.5449 | -35.1013
24.6978 | -35.5453 | -35.1073
24.7039 | -35.5457 | -35.1132
24.71 | -35.546 | -35.1195
24.7162 | -35.5465 | -35.1254
24.7223 | -35.5469 | -35.1312
24.7284 | -35.5473 | -35.1362
24.7345 | -35.5477 | -35.142
24.7406 | -35.5481 | -35.148
24.7467 | -35.5485 | -35.154
24.7528 | -35.5489 | -35.16
24.7589 | -35.5493 | -35.1658
24.765 | -35.5497 | -35.1718
24.7711 | -35.5501 | -35.1778
24.7772 | -35.5504 | -35.1838
24.7833 | -35.5508 | -35.1898
24.7894 | -35.5511 | -35.1961
24.7955 | -35.5515 | -35.2018
24.8016 | -35.5518 | -35.2074
24.8077 | -35.5522 | -35.2128
24.8138 | -35.5526 | -35.2186
24.8199 | -35.5529 | -35.2241
24.826 | -35.5533 | -35.2301
24.8321 | -35.5537 | -35.2356
24.8382 | -35.5541 | -35.2412
24.8443 | -35.5545 | -35.2468
24.8504 | -35.5548 | -35.2522
24.8565 | -35.5552 | -35.258
24.8627 | -35.5556 | -35.2639
24.8688 | -35.5559 | -35.2693
24.8749 | -35.5563 | -35.2754
24.881 | -35.5567 | -35.2814
24.8871 | -35.557 | -35.2873
24.8932 | -35.5574 | -35.293
24.8993 | -35.5577 | -35.2987
24.9054 | -35.5581 | -35.3044
24.9115 | -35.5585 | -35.3101
24.9176 | -35.5589 | -35.316
24.9237 | -35.5594 | -35.3217
24.9298 | -35.5597 | -35.3276
24.9359 | -35.5601 | -35.3335
24.942 | -35.5604 | -35.339
24.9481 | -35.5608 | -35.3449
24.9542 | -35.5611 | -35.3506
24.9603 | -35.5615 | -35.3565
24.9664 | -35.5619 | -35.3619
24.9725 | -35.5622 | -35.3674
24.9786 | -35.5627 | -35.3729
24.9847 | -35.5631 | -35.3785
24.9908 | -35.5635 | -35.3843
24.9969 | -35.5638 | -35.3901
25.0031 | -35.5642 | -35.3961
25.0092 | -35.5646 | -35.4015
25.0153 | -35.5649 | -35.4072
25.0214 | -35.5652 | -35.4126
25.0275 | -35.5656 | -35.4185
25.0336 | -35.566 | -35.4238
25.0397 | -35.5664 | -35.4298
25.0458 | -35.5668 | -35.4356
25.0519 | -35.5671 | -35.4414
25.058 | -35.5675 | -35.447
25.0641 | -35.5679 | -35.4527
25.0702 | -35.5682 | -35.4584
25.0763 | -35.5686 | -35.464
25.0824 | -35.569 | -35.4694
25.0885 | -35.5693 | -35.4751
25.0946 | -35.5697 | -35.4806
25.1007 | -35.5701 | -35.4866
25.1068 | -35.5705 | -35.4924
25.1129 | -35.5709 | -35.4983
25.119 | -35.5714 | -35.5037
25.1251 | -35.5717 | -35.5095
25.1312 | -35.5721 | -35.5151
25.1373 | -35.5725 | -35.5209
25.1435 | -35.5729 | -35.5265
25.1496 | -35.5733 | -35.5325
25.1557 | -35.5737 | -35.5383
25.1618 | -35.574 | -35.5444
25.1679 | -35.5744 | -35.55
25.174 | -35.5748 | -35.5559
25.1801 | -35.5752 | -35.562
25.1862 | -35.5755 | -35.5679
25.1923 | -35.5758 | -35.5736
25.1984 | -35.5762 | -35.5795
25.2045 | -35.5766 | -35.585
25.2106 | -35.577 | -35.5906
25.2167 | -35.5774 | -35.5963
25.2228 | -35.5777 | -35.602
25.2289 | -35.5781 | -35.6076
25.235 | -35.5785 | -35.6135
25.2411 | -35.5789 | -35.619
25.2472 | -35.5792 | -35.6246
25.2533 | -35.5796 | -35.6301
25.2594 | -35.58 | -35.636
25.2655 | -35.5804 | -35.6417
25.2716 | -35.5808 | -35.6478
25.2777 | -35.5811 | -35.6535
25.2838 | -35.5815 | -35.6593
25.29 | -35.5819 | -35.6647
25.2961 | -35.5824 | -35.6706
25.3022 | -35.5828 | -35.6763
25.3083 | -35.5832 | -35.6819
25.3144 | -35.5835 | -35.6875
25.3205 | -35.5839 | -35.6935
25.3266 | -35.5842 | -35.6989
25.3327 | -35.5845 | -35.7043
25.3388 | -35.5849 | -35.7102
25.3449 | -35.5852 | -35.716
25.351 | -35.5856 | -35.7218
25.3571 | -35.586 | -35.7276
25.3632 | -35.5864 | -35.733
25.3693 | -35.5867 | -35.7384
25.3754 | -35.5871 | -35.7439
25.3815 | -35.5875 | -35.7495
25.3876 | -35.5879 | -35.7552
25.3937 | -35.5883 | -35.7608
25.3998 | -35.5886 | -35.7659
25.4059 | -35.589 | -35.7719
25.412 | -35.5894 | -35.7774
25.4181 | -35.5897 | -35.7832
25.4242 | -35.5901 | -35.7888
25.4304 | -35.5904 | -35.7944
25.4365 | -35.5908 | -35.8004
25.4426 | -35.5911 | -35.8059
25.4487 | -35.5914 | -35.8114
25.4548 | -35.5918 | -35.8173
25.4609 | -35.5921 | -35.8229
25.467 | -35.5926 | -35.8285
25.4731 | -35.593 | -35.8341
25.4792 | -35.5933 | -35.8397
25.4853 | -35.5937 | -35.8454
25.4914 | -35.5942 | -35.851
25.4975 | -35.5945 | -35.8567
25.5036 | -35.5949 | -35.8621
25.5097 | -35.5953 | -35.8677
25.5158 | -35.5957 | -35.8735
25.5219 | -35.5961 | -35.8793
25.528 | -35.5964 | -35.8849
25.5341 | -35.5968 | -35.8907
25.5402 | -35.5972 | -35.8965
25.5463 | -35.5975 | -35.9024
25.5524 | -35.5979 | -35.9082
25.5585 | -35.5983 | -35.9139
25.5646 | -35.5987 | -35.9196
25.5707 | -35.5991 | -35.9255
25.5769 | -35.5994 | -35.9303
25.583 | -35.5998 | -35.9364
25.5891 | -35.6001 | -35.942
25.5952 | -35.6005 | -35.9479
25.6013 | -35.6009 | -35.9534
25.6074 | -35.6013 | -35.9588
25.6135 | -35.6016 | -35.9645
25.6196 | -35.602 | -35.9701
25.6257 | -35.6023 | -35.9757
25.6318 | -35.6027 | -35.9815
25.6379 | -35.6031 | -35.9876
25.644 | -35.6034 | -35.9928
25.6501 | -35.6037 | -35.9981
25.6562 | -35.6041 | -36.0036
25.6623 | -35.6044 | -36.009
25.6684 | -35.6048 | -36.0147
25.6745 | -35.6052 | -36.0201
25.6806 | -35.6055 | -36.0255
25.6867 | -35.6059 | -36.031
25.6928 | -35.6062 | -36.0364
25.6989 | -35.6066 | -36.0419
25.705 | -35.607 | -36.0481
25.7111 | -35.6073 | -36.0536
25.7173 | -35.6077 | -36.0588
25.7234 | -35.6081 | -36.0644
25.7295 | -35.6085 | -36.0696
25.7356 | -35.6089 | -36.0752
25.7417 | -35.6092 | -36.0809
25.7478 | -35.6096 | -36.0868
25.7539 | -35.6099 | -36.0923
25.76 | -35.6102 | -36.0976
25.7661 | -35.6106 | -36.1031
25.7722 | -35.6109 | -36.1089
25.7783 | -35.6113 | -36.1144
25.7844 | -35.6117 | -36.1198
25.7905 | -35.612 | -36.1253
25.7966 | -35.6124 | -36.131
25.8027 | -35.6127 | -36.1365
25.8088 | -35.613 | -36.142
25.8149 | -35.6134 | -36.1477
25.821 | -35.6137 | -36.1532
25.8271 | -35.614 | -36.1588
25.8332 | -35.6144 | -36.1645
25.8393 | -35.6148 | -36.1701
25.8454 | -35.6151 | -36.1758
25.8515 | -35.6155 | -36.1816
25.8576 | -35.6158 | -36.1873
25.8638 | -35.6162 | -36.1931
25.8699 | -35.6166 | -36.199
25.876 | -35.617 | -36.2047
25.8821 | -35.6173 | -36.2102
25.8882 | -35.6177 | -36.2156
25.8943 | -35.618 | -36.2212
25.9004 | -35.6184 | -36.2266
25.9065 | -35.6188 | -36.2324
25.9126 | -35.6191 | -36.2381
25.9187 | -35.6195 | -36.2439
25.9248 | -35.6198 | -36.2491
25.9309 | -35.6202 | -36.2547
25.937 | -35.6206 | -36.2602
25.9431 | -35.6209 | -36.2657
25.9492 | -35.6213 | -36.271
25.9553 | -35.6216 | -36.2766
25.9614 | -35.622 | -36.2822
25.9675 | -35.6224 | -36.2875
25.9736 | -35.6227 | -36.2933
25.9797 | -35.6231 | -36.2988
25.9858 | -35.6234 | -36.3042
25.9919 | -35.6238 | -36.3098
25.998 | -35.6242 | -36.3153
26.0042 | -35.6245 | -36.3208
26.0103 | -35.6248 | -36.3267
26.0164 | -35.6251 | -36.3319
26.0225 | -35.6255 | -36.3371
26.0286 | -35.6258 | -36.3427
26.0347 | -35.6261 | -36.348
26.0408 | -35.6265 | -36.3532
26.0469 | -35.6268 | -36.3586
26.053 | -35.6272 | -36.3644
26.0591 | -35.6276 | -36.3697
26.0652 | -35.6278 | -36.3752
26.0713 | -35.6282 | -36.3804
26.0774 | -35.6285 | -36.3859
26.0835 | -35.6289 | -36.3918
26.0896 | -35.6293 | -36.3978
26.0957 | -35.6297 | -36.4032
26.1018 | -35.6301 | -36.4083
26.1079 | -35.6304 | -36.4141
26.114 | -35.6307 | -36.4195
26.1201 | -35.631 | -36.4256
26.1262 | -35.6314 | -36.4308
26.1323 | -35.6318 | -36.4361
26.1384 | -35.6322 | -36.4414
26.1445 | -35.6326 | -36.4471
26.1507 | -35.6329 | -36.4527
26.1568 | -35.6333 | -36.4584
26.1629 | -35.6336 | -36.4641
26.169 | -35.634 | -36.4692
26.1751 | -35.6344 | -36.4745
26.1812 | -35.6347 | -36.4801
26.1873 | -35.6351 | -36.4854
26.1934 | -35.6355 | -36.491
26.1995 | -35.6358 | -36.4967
26.2056 | -35.6361 | -36.5018
26.2117 | -35.6364 | -36.5075
26.2178 | -35.6368 | -36.5131
26.2239 | -35.6371 | -36.5188
26.23 | -35.6374 | -36.5243
26.2361 | -35.6377 | -36.5297
26.2422 | -35.638 | -36.5356
26.2483 | -35.6384 | -36.5412
26.2544 | -35.6387 | -36.5466
26.2605 | -35.6391 | -36.5516
26.2666 | -35.6395 | -36.5575
26.2727 | -35.6398 | -36.5635
26.2788 | -35.6401 | -36.5687
26.2849 | -35.6405 | -36.5742
26.2911 | -35.6409 | -36.5802
26.2972 | -35.6412 | -36.5858
26.3033 | -35.6416 | -36.5912
26.3094 | -35.6419 | -36.5965
26.3155 | -35.6422 | -36.6019
26.3216 | -35.6425 | -36.6077
26.3277 | -35.6429 | -36.6131
26.3338 | -35.6433 | -36.6188
26.3399 | -35.6437 | -36.6247
26.346 | -35.644 | -36.6298
26.3521 | -35.6444 | -36.6349
26.3582 | -35.6448 | -36.6401
26.3643 | -35.6451 | -36.6455
26.3704 | -35.6454 | -36.6507
26.3765 | -35.6458 | -36.6562
26.3826 | -35.6461 | -36.6616
26.3887 | -35.6465 | -36.6672
26.3948 | -35.6469 | -36.6729
26.4009 | -35.6472 | -36.6784
26.407 | -35.6476 | -36.6839
26.4131 | -35.6479 | -36.6892
26.4192 | -35.6483 | -36.6943
26.4253 | -35.6487 | -36.6997
26.4314 | -35.649 | -36.705
26.4376 | -35.6493 | -36.711
26.4437 | -35.6497 | -36.7164
26.4498 | -35.65 | -36.7216
26.4559 | -35.6504 | -36.7269
26.462 | -35.6508 | -36.7326
26.4681 | -35.6511 | -36.7378
26.4742 | -35.6514 | -36.743
26.4803 | -35.6518 | -36.7485
26.4864 | -35.6521 | -36.7537
26.4925 | -35.6524 | -36.7588
26.4986 | -35.6527 | -36.7643
26.5047 | -35.6531 | -36.7694
26.5108 | -35.6534 | -36.7749
26.5169 | -35.6538 | -36.7802
26.523 | -35.6541 | -36.7853
26.5291 | -35.6544 | -36.7904
26.5352 | -35.6548 | -36.796
26.5413 | -35.6551 | -36.8016
26.5474 | -35.6555 | -36.8072
26.5535 | -35.6558 | -36.8129
26.5596 | -35.6561 | -36.8182
26.5657 | -35.6564 | -36.8236
26.5718 | -35.6568 | -36.8291
26.578 | -35.6571 | -36.8347
26.5841 | -35.6575 | -36.8405
26.5902 | -35.6579 | -36.8463
26.5963 | -35.6582 | -36.8512
26.6024 | -35.6585 | -36.8565
26.6085 | -35.6589 | -36.8619
26.6146 | -35.6592 | -36.8678
26.6207 | -35.6595 | -36.8738
26.6268 | -35.6599 | -36.879
26.6329 | -35.6603 | -36.8842
26.639 | -35.6606 | -36.8893
26.6451 | -35.6609 | -36.8948
26.6512 | -35.6612 | -36.8996
26.6573 | -35.6616 | -36.9048
26.6634 | -35.662 | -36.9094
26.6695 | -35.6623 | -36.9152
26.6756 | -35.6628 | -36.9206
26.6817 | -35.663 | -36.9258
26.6878 | -35.6633 | -36.9312
26.6939 | -35.6637 | -36.9365
26.7 | -35.664 | -36.9417
26.7061 | -35.6644 | -36.9468
26.7122 | -35.6648 | -36.9521
26.7183 | -35.6651 | -36.9577
26.7245 | -35.6655 | -36.9632
26.7306 | -35.6658 | -36.9687
26.7367 | -35.6662 | -36.9733
26.7428 | -35.6665 | -36.9785
26.7489 | -35.6669 | -36.9838
26.755 | -35.6673 | -36.9893
26.7611 | -35.6676 | -36.9949
26.7672 | -35.6679 | -37.0004
26.7733 | -35.6682 | -37.0058
26.7794 | -35.6686 | -37.011
26.7855 | -35.6689 | -37.0162
26.7916 | -35.6693 | -37.0218
26.7977 | -35.6696 | -37.0273
26.8038 | -35.6699 | -37.0327
26.8099 | -35.6703 | -37.0378
26.816 | -35.6707 | -37.0431
26.8221 | -35.671 | -37.0484
26.8282 | -35.6714 | -37.0535
26.8343 | -35.6717 | -37.0592
26.8404 | -35.672 | -37.064
26.8465 | -35.6723 | -37.0691
26.8526 | -35.6727 | -37.0749
26.8587 | -35.673 | -37.08
26.8649 | -35.6733 | -37.0852
26.871 | -35.6737 | -37.0902
26.8771 | -35.674 | -37.0955
26.8832 | -35.6743 | -37.1006
26.8893 | -35.6747 | -37.1059
26.8954 | -35.6751 | -37.1111
26.9015 | -35.6754 | -37.1165
26.9076 | -35.6757 | -37.1214
26.9137 | -35.676 | -37.1268
26.9198 | -35.6764 | -37.1321
26.9259 | -35.6767 | -37.1375
26.932 | -35.677 | -37.143
26.9381 | -35.6774 | -37.1484
26.9442 | -35.6778 | -37.1536
26.9503 | -35.6781 | -37.1589
26.9564 | -35.6785 | -37.1642
26.9625 | -35.6788 | -37.1698
26.9686 | -35.6791 | -37.1754
26.9747 | -35.6795 | -37.1805
26.9808 | -35.6799 | -37.1854
26.9869 | -35.6802 | -37.1908
26.993 | -35.6806 | -37.1961
26.9991 | -35.6809 | -37.2015
27.0052 | -35.6813 | -37.2074
27.0114 | -35.6816 | -37.2125
27.0175 | -35.6819 | -37.2183
27.0236 | -35.6823 | -37.2236
27.0297 | -35.6826 | -37.2289
27.0358 | -35.6829 | -37.2343
27.0419 | -35.6833 | -37.2398
27.048 | -35.6836 | -37.245
27.0541 | -35.684 | -37.25
27.0602 | -35.6844 | -37.2554
27.0663 | -35.6848 | -37.2608
27.0724 | -35.6851 | -37.2658
27.0785 | -35.6855 | -37.2711
27.0846 | -35.6859 | -37.2761
27.0907 | -35.6862 | -37.2816
27.0968 | -35.6866 | -37.2871
27.1029 | -35.6869 | -37.2925
27.109 | -35.6874 | -37.2985
27.1151 | -35.6876 | -37.3039
27.1212 | -35.6879 | -37.3089
27.1273 | -35.6883 | -37.314
27.1334 | -35.6886 | -37.3188
27.1395 | -35.689 | -37.3244
27.1456 | -35.6893 | -37.3299
27.1518 | -35.6897 | -37.3355
27.1579 | -35.69 | -37.3409
27.164 | -35.6903 | -37.3456
27.1701 | -35.6907 | -37.3509
27.1762 | -35.691 | -37.3561
27.1823 | -35.6913 | -37.3614
27.1884 | -35.6917 | -37.3668
27.1945 | -35.692 | -37.3719
27.2006 | -35.6923 | -37.3773
27.2067 | -35.6926 | -37.3824
27.2128 | -35.693 | -37.3882
27.2189 | -35.6932 | -37.3932
27.225 | -35.6936 | -37.3989
27.2311 | -35.6939 | -37.4043
27.2372 | -35.6943 | -37.41
27.2433 | -35.6946 | -37.4152
27.2494 | -35.695 | -37.4207
27.2555 | -35.6954 | -37.426
27.2616 | -35.6958 | -37.4311
27.2677 | -35.6961 | -37.4364
27.2738 | -35.6964 | -37.4417
27.2799 | -35.6967 | -37.4469
27.286 | -35.697 | -37.4524
27.2921 | -35.6974 | -37.4581
27.2983 | -35.6978 | -37.4632
27.3044 | -35.6981 | -37.4687
27.3105 | -35.6984 | -37.4744
27.3166 | -35.6987 | -37.4797
27.3227 | -35.6991 | -37.485
27.3288 | -35.6994 | -37.49
27.3349 | -35.6997 | -37.495
27.341 | -35.6999 | -37.5001
27.3471 | -35.7003 | -37.505
27.3532 | -35.7007 | -37.5105
27.3593 | -35.701 | -37.5161
27.3654 | -35.7014 | -37.5217
27.3715 | -35.7018 | -37.527
27.3776 | -35.7021 | -37.5324
27.3837 | -35.7024 | -37.5373
27.3898 | -35.7029 | -37.5425
27.3959 | -35.7032 | -37.548
27.402 | -35.7035 | -37.5529
27.4081 | -35.7038 | -37.5585
27.4142 | -35.7041 | -37.5634
27.4203 | -35.7045 | -37.5688
27.4264 | -35.7048 | -37.5745
27.4325 | -35.7051 | -37.5795
27.4387 | -35.7055 | -37.5843
27.4448 | -35.7058 | -37.5897
27.4509 | -35.7061 | -37.5954
27.457 | -35.7065 | -37.6009
27.4631 | -35.7068 | -37.6062
27.4692 | -35.7072 | -37.6115
27.4753 | -35.7075 | -37.6165
27.4814 | -35.7078 | -37.6221
27.4875 | -35.7081 | -37.627
27.4936 | -35.7085 | -37.632
27.4997 | -35.7089 | -37.6372
27.5058 | -35.7092 | -37.6424
27.5119 | -35.7095 | -37.6481
27.518 | -35.7099 | -37.6533
27.5241 | -35.7102 | -37.6581
27.5302 | -35.7106 | -37.6634
27.5363 | -35.7109 | -37.6685
27.5424 | -35.7112 | -37.6734
27.5485 | -35.7115 | -37.6786
27.5546 | -35.7119 | -37.684
27.5607 | -35.7122 | -37.6895
27.5668 | -35.7125 | -37.6949
27.5729 | -35.7128 | -37.6997
27.5791 | -35.7132 | -37.7049
27.5852 | -35.7135 | -37.7101
27.5913 | -35.7138 | -37.7149
27.5974 | -35.7142 | -37.7202
27.6035 | -35.7145 | -37.7257
27.6096 | -35.7148 | -37.7306
27.6157 | -35.7152 | -37.7354
27.6218 | -35.7155 | -37.7408
27.6279 | -35.7158 | -37.7461
27.634 | -35.7161 | -37.7515
27.6401 | -35.7164 | -37.757
27.6462 | -35.7167 | -37.7622
27.6523 | -35.7171 | -37.7671
27.6584 | -35.7174 | -37.7721
27.6645 | -35.7177 | -37.778
27.6706 | -35.718 | -37.7835
27.6767 | -35.7184 | -37.7883
27.6828 | -35.7187 | -37.7934
27.6889 | -35.7191 | -37.7994
27.695 | -35.7194 | -37.8044
27.7011 | -35.7197 | -37.8097
27.7072 | -35.7201 | -37.8145
27.7133 | -35.7204 | -37.8195
27.7194 | -35.7207 | -37.825
27.7256 | -35.721 | -37.8305
27.7317 | -35.7213 | -37.8359
27.7378 | -35.7216 | -37.8408
27.7439 | -35.7219 | -37.8462
27.75 | -35.7223 | -37.8516
27.7561 | -35.7225 | -37.8564
27.7622 | -35.7229 | -37.8616
27.7683 | -35.7233 | -37.8668
27.7744 | -35.7237 | -37.872
27.7805 | -35.724 | -37.8768
27.7866 | -35.7243 | -37.8823
27.7927 | -35.7246 | -37.887
27.7988 | -35.7249 | -37.8923
27.8049 | -35.7253 | -37.8976
27.811 | -35.7256 | -37.9035
27.8171 | -35.7259 | -37.9085
27.8232 | -35.7262 | -37.9133
27.8293 | -35.7266 | -37.9187
27.8354 | -35.727 | -37.924
27.8415 | -35.7274 | -37.929
27.8476 | -35.7277 | -37.9343
27.8537 | -35.7281 | -37.9393
27.8598 | -35.7285 | -37.9443
27.866 | -35.7288 | -37.9494
27.8721 | -35.7291 | -37.9546
27.8782 | -35.7294 | -37.9597
27.8843 | -35.7297 | -37.965
27.8904 | -35.73 | -37.9705
27.8965 | -35.7303 | -37.9752
27.9026 | -35.7307 | -37.9807
27.9087 | -35.7311 | -37.9858
27.9148 | -35.7315 | -37.9911
27.9209 | -35.7318 | -37.9964
27.927 | -35.732 | -38.0017
27.9331 | -35.7324 | -38.0064
27.9392 | -35.7328 | -38.0115
27.9453 | -35.7331 | -38.0163
27.9514 | -35.7334 | -38.0217
27.9575 | -35.7337 | -38.0269
27.9636 | -35.7341 | -38.0315
27.9697 | -35.7344 | -38.0368
27.9758 | -35.7347 | -38.0422
27.9819 | -35.7351 | -38.0475
27.988 | -35.7354 | -38.0526
27.9941 | -35.7356 | -38.0577
28.0002 | -35.7359 | -38.0634
28.0063 | -35.7362 | -38.0692
28.0125 | -35.7365 | -38.0735
28.0186 | -35.7369 | -38.0787
28.0247 | -35.7373 | -38.0844
28.0308 | -35.7376 | -38.0892
28.0369 | -35.7379 | -38.0943
28.043 | -35.7382 | -38.0994
28.0491 | -35.7385 | -38.1046
28.0552 | -35.7388 | -38.1096
28.0613 | -35.7391 | -38.1149
28.0674 | -35.7395 | -38.1198
28.0735 | -35.7398 | -38.125

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4222 | -42.6388
21.4908 | -35.4856 | -35.3277
16.4948 | -34.866 | -27.9294
12.6602 | -34.4957 | -21.6161
9.7171 | -34.2769 | -16.6789
7.4582 | -34.1291 | -12.6663
5.7244 | -34.0187 | -9.4189
4.3936 | -33.9458 | -6.7955
3.3722 | -33.9318 | -4.7276
2.5883 | -33.9353 | -3.1585
1.9866 | -33.9299 | -1.9062
1.5248 | -33.93 | -0.90261
1.1703 | -33.9377 | -0.24636
1 | -33.9417 | 0.083944
0.89822 | -33.9334 | 0.35633
0.68941 | -33.9344 | 0.94135
0.52919 | -33.9297 | 1.3949
0.40612 | -33.9343 | 2.0005
0.31171 | -33.914 | 2.4586
0.23928 | -33.9355 | 3.1724
0.18358 | -33.9151 | 3.6048
0.14091 | -33.9535 | 4.4706
0.10819 | -33.9514 | 5.6261
0.083017 | -33.9947 | 7.099
0.063705 | -34.0502 | 8.9719
0.048923 | -34.1075 | 11.4612
0.037479 | -34.2053 | 14.64
0.028801 | -34.3812 | 18.6629
0.022078 | -34.6978 | 24.0529
0.017023 | -35.1808 | 30.4809
0.013018 | -35.9827 | 38.1527
0.010014 | -37.0988 | 46.0644

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003666 | 1.9014
0.007331 | -28.7798
0.010997 | -35.9049
0.014662 | -41.4345
0.018328 | -43.8841
0.021993 | -45.6673
0.025659 | -46.8693
0.029324 | -49.227
0.03299 | -51.3073
0.036655 | -50.6978
0.040321 | -51.3119
0.043986 | -52.6917
0.047652 | -54.5622
0.051317 | -56.4196
0.054983 | -56.813
0.058648 | -56.2837
0.062314 | -56.6785
0.065979 | -57.9841
0.069645 | -59.7861
0.07331 | -58.8365
0.076976 | -59.8926
0.080641 | -61.3564
0.084307 | -62.1029
0.087973 | -62.3153
0.091638 | -61.6045
0.095304 | -62.5146
0.098969 | -63.2502
0.10264 | -63.8089
0.1063 | -64.0159
0.10997 | -63.8709
0.11363 | -64.923
0.1173 | -66.4104
0.12096 | -67.2286
0.12463 | -66.3376
0.12829 | -66.3833
0.13196 | -67.4714
0.13562 | -67.5365
0.13929 | -67.7731
0.14295 | -67.586
0.14662 | -67.0816
0.15029 | -68.5929
0.15395 | -67.9311
0.15762 | -68.9532
0.16128 | -69.8295
0.16495 | -70.3753
0.16861 | -70.0166
0.17228 | -69.7374
0.17594 | -69.1724
0.17961 | -70.2799
0.18328 | -71.1484
0.18694 | -70.5758
0.19061 | -70.0378
0.19427 | -70.357
0.19794 | -70.4671
0.2016 | -71.6237
0.20527 | -72.1734
0.20894 | -72.6845
0.2126 | -71.5999
0.21627 | -71.5855
0.21993 | -73.1958
0.2236 | -72.9335
0.22726 | -72.4687
0.23093 | -72.3954
0.23459 | -73.3307
0.23826 | -73.5653
0.24192 | -73.7675
0.24559 | -73.9089
0.24926 | -73.9013
0.25292 | -74.3781
0.25659 | -74.0826
0.26025 | -73.298
0.26392 | -74.2998
0.26758 | -75.1168
0.27125 | -74.6444
0.27491 | -75.1579
0.27858 | -75.474
0.28225 | -74.4186
0.28591 | -74.3973
0.28958 | -74.684
0.29324 | -74.007
0.29691 | -75.4048
0.30057 | -76.0146
0.30424 | -76.3315
0.3079 | -76.6198
0.31157 | -76.0134
0.31523 | -76.2067
0.3189 | -76.1198
0.32257 | -76.2796
0.32623 | -76.2161
0.3299 | -76.6355
0.33356 | -76.2325
0.33723 | -76.1208
0.34089 | -76.6926
0.34456 | -77.7645
0.34823 | -77.9701
0.35189 | -77.1617
0.35556 | -77.3024
0.35922 | -77.5083
0.36289 | -76.9503
0.36655 | -78.1547
0.37022 | -78.521
0.37388 | -77.6259
0.37755 | -77.4839
0.38121 | -77.3698
0.38488 | -77.024
0.38854 | -77.8814
0.39221 | -78.4316
0.39588 | -78.1788
0.39954 | -77.9103
0.40321 | -78.023
0.40687 | -78.4959
0.41054 | -78.3303
0.4142 | -77.9596
0.41787 | -79.091
0.42153 | -80.0298
0.4252 | -79.9343
0.42887 | -79.3887
0.43253 | -78.9296
0.4362 | -80.0128
0.43986 | -80.6689
0.44353 | -78.9364
0.44719 | -79.1255
0.45086 | -79.6785
0.45453 | -80.5665
0.45819 | -79.9809
0.46186 | -79.3923
0.46552 | -81.0115
0.46919 | -81.2737
0.47285 | -80.493
0.47652 | -79.8186
0.48018 | -80.1616
0.48385 | -80.1126
0.48751 | -80.251
0.49118 | -79.3993
0.49484 | -79.4259
0.49851 | -80.808
0.50218 | -80.3824
0.50584 | -80.0168
0.50951 | -81.5666
0.51317 | -81.9722
0.51684 | -81.0782
0.5205 | -81.424
0.52417 | -81.3066
0.52784 | -80.8018
0.5315 | -80.7284
0.53517 | -80.901
0.53883 | -81.3783
0.5425 | -81.5679
0.54616 | -81.3542
0.54983 | -81.2149
0.55349 | -81.5034
0.55716 | -81.978
0.56083 | -82.3456
0.56449 | -82.0579
0.56816 | -81.1933
0.57182 | -80.7931
0.57549 | -81.4297
0.57915 | -82.2997
0.58282 | -82.6809
0.58648 | -82.5236
0.59015 | -82.2854
0.59381 | -82.2937
0.59748 | -82.4545
0.60115 | -82.5543
0.60481 | -82.6197
0.60848 | -82.5253
0.61214 | -82.3643
0.61581 | -82.0952
0.61947 | -81.9509
0.62314 | -82.2669
0.6268 | -82.951
0.63047 | -83.1592
0.63414 | -83.261
0.6378 | -83.6384
0.64147 | -83.7605
0.64513 | -83.401
0.6488 | -83.1849
0.65246 | -83.5493
0.65613 | -83.6904
0.65979 | -83.6849
0.66346 | -83.6511
0.66712 | -83.5652
0.67079 | -83.3984
0.67446 | -83.3988
0.67812 | -83.4101
0.68179 | -83.2786
0.68545 | -83.2097
0.68912 | -83.3512
0.69278 | -83.6005
0.69645 | -83.761
0.70012 | -83.538
0.70378 | -83.6629
0.70745 | -84.3992
0.71111 | -84.7722
0.71478 | -84.7131
0.71844 | -84.8596
0.72211 | -84.8926
0.72577 | -85.1938
0.72944 | -85.9281
0.7331 | -86.542
0.73677 | -86.1849
0.74043 | -84.7835
0.7441 | -84.0502
0.74777 | -84.0445
0.75143 | -84.1064
0.7551 | -84.231
0.75876 | -84.5626
0.76243 | -84.9641
0.76609 | -85.5151
0.76976 | -85.6176
0.77343 | -85.2631
0.77709 | -85.3729
0.78076 | -85.5156
0.78442 | -85.7753
0.78809 | -86.0257
0.79175 | -85.7066
0.79542 | -85.2121
0.79908 | -84.7882
0.80275 | -84.7239
0.80641 | -85.1979
0.81008 | -85.5043
0.81375 | -85.2202
0.81741 | -84.9176
0.82108 | -85.1098
0.82474 | -85.5103
0.82841 | -85.8716
0.83207 | -86.0875
0.83574 | -85.9258
0.8394 | -85.8156
0.84307 | -85.962
0.84674 | -86.1339
0.8504 | -86.2841
0.85407 | -86.7708
0.85773 | -87.0434
0.8614 | -86.704
0.86506 | -86.3498
0.86873 | -86.2144
0.87239 | -86.3735
0.87606 | -86.7065
0.87972 | -86.9461
0.88339 | -87.1792
0.88706 | -87.1463
0.89072 | -86.7002
0.89439 | -86.3095
0.89805 | -85.9189
0.90172 | -85.7221
0.90538 | -85.9069
0.90905 | -86.1705
0.91272 | -86.2891
0.91638 | -86.3868
0.92005 | -86.4802
0.92371 | -86.5421
0.92738 | -86.8143
0.93104 | -87.3253
0.93471 | -87.3688
0.93837 | -86.9273
0.94204 | -86.6771
0.9457 | -86.4711
0.94937 | -86.3321
0.95304 | -86.566
0.9567 | -87.039
0.96037 | -87.2786
0.96403 | -87.2916
0.9677 | -87.0613
0.97136 | -86.7239
0.97503 | -86.3047
0.97869 | -86.295
0.98236 | -86.9394
0.98603 | -87.5627
0.98969 | -87.8878
0.99336 | -87.8704
0.99702 | -87.3711
1.0007 | -87.3972
1.0044 | -87.7969
1.008 | -87.7373
1.0117 | -87.7649
1.0153 | -87.9026
1.019 | -88.0438
1.0227 | -88.261
1.0263 | -88.8628
1.03 | -89.2931
1.0337 | -89.0719
1.0373 | -88.5573
1.041 | -88.1344
1.0447 | -88.3344
1.0483 | -88.7423
1.052 | -88.9837
1.0557 | -88.7277
1.0593 | -88.3561
1.063 | -88.0896
1.0667 | -88.165
1.0703 | -88.406
1.074 | -88.4667
1.0777 | -88.2603
1.0813 | -87.9659
1.085 | -87.7407
1.0887 | -87.6994
1.0923 | -87.7885
1.096 | -87.7247
1.0997 | -87.5366
1.1033 | -87.6178
1.107 | -88.0619
1.1107 | -88.5015
1.1143 | -88.584
1.118 | -88.4909
1.1216 | -88.5446
1.1253 | -88.7055
1.129 | -88.7122
1.1326 | -88.4474
1.1363 | -88.1186
1.14 | -88.0814
1.1436 | -88.4107
1.1473 | -88.8038
1.151 | -88.9658
1.1546 | -88.9236
1.1583 | -88.8083
1.162 | -88.6819
1.1656 | -88.6377
1.1693 | -88.7957
1.173 | -88.9954
1.1766 | -88.9412
1.1803 | -88.7469
1.184 | -88.6487
1.1876 | -88.6088
1.1913 | -88.6026
1.195 | -88.7093
1.1986 | -88.9282
1.2023 | -89.0745
1.206 | -88.8853
1.2096 | -88.4282
1.2133 | -88.1147
1.217 | -88.2113
1.2206 | -88.6298
1.2243 | -89.073
1.228 | -89.3156
1.2316 | -89.4307
1.2353 | -89.5321
1.2389 | -89.5499
1.2426 | -89.4982
1.2463 | -89.4701
1.2499 | -89.4851
1.2536 | -89.5292
1.2573 | -89.4819
1.2609 | -89.2561
1.2646 | -88.9802
1.2683 | -88.7987
1.2719 | -88.8099
1.2756 | -89.0338
1.2793 | -89.3513
1.2829 | -89.6478
1.2866 | -89.902
1.2903 | -90.0468
1.2939 | -90.0091
1.2976 | -89.8245
1.3013 | -89.6585
1.3049 | -89.6492
1.3086 | -89.7815
1.3123 | -89.8706
1.3159 | -89.8475
1.3196 | -89.8286
1.3233 | -89.9008
1.3269 | -90.0294
1.3306 | -90.0835
1.3342 | -90.0821
1.3379 | -90.083
1.3416 | -89.9778
1.3452 | -89.9249
1.3489 | -90.1632
1.3526 | -90.3808
1.3562 | -90.2609
1.3599 | -90.0847
1.3636 | -90.0346
1.3672 | -90.0364
1.3709 | -90.0729
1.3746 | -90.0901
1.3782 | -90.0094
1.3819 | -89.8988
1.3856 | -89.9049
1.3892 | -90.0261
1.3929 | -90.1618
1.3966 | -90.2869
1.4002 | -90.4528
1.4039 | -90.7128
1.4076 | -90.9966
1.4112 | -91.1872
1.4149 | -91.23
1.4186 | -91.0833
1.4222 | -90.8724
1.4259 | -90.8114
1.4296 | -90.7916
1.4332 | -90.6198
1.4369 | -90.3393
1.4405 | -90.1056
1.4442 | -90.0795
1.4479 | -90.3188
1.4515 | -90.5946
1.4552 | -90.5607
1.4589 | -90.2875
1.4625 | -90.2167
1.4662 | -90.5053
1.4699 | -90.9144
1.4735 | -91.1822
1.4772 | -91.3607
1.4809 | -91.6344
1.4845 | -91.8398
1.4882 | -91.6599
1.4919 | -91.2878
1.4955 | -91.0009
1.4992 | -90.8083
1.5029 | -90.7477
1.5065 | -90.8866
1.5102 | -91.1177
1.5139 | -91.1719
1.5175 | -90.9897
1.5212 | -90.9033
1.5249 | -91.1334
1.5285 | -91.5597
1.5322 | -91.897
1.5359 | -91.8859
1.5395 | -91.4857
1.5432 | -90.977
1.5469 | -90.687
1.5505 | -90.6541
1.5542 | -90.7358
1.5578 | -90.9061
1.5615 | -91.1298
1.5652 | -91.254
1.5688 | -91.2383
1.5725 | -91.1993
1.5762 | -91.3214
1.5798 | -91.6625
1.5835 | -91.718
1.5872 | -91.694
1.5908 | -91.6502
1.5945 | -91.7515
1.5982 | -91.9899
1.6018 | -92.2393
1.6055 | -92.4143
1.6092 | -92.5497
1.6128 | -92.6812
1.6165 | -92.7553
1.6202 | -92.6581
1.6238 | -92.3781
1.6275 | -92.0947
1.6312 | -91.966
1.6348 | -92.0076
1.6385 | -92.123
1.6422 | -92.1919
1.6458 | -92.2047
1.6495 | -92.2412
1.6531 | -92.3084
1.6568 | -92.3121
1.6605 | -92.1985
1.6641 | -91.9802
1.6678 | -91.7001
1.6715 | -91.4933
1.6751 | -91.4908
1.6788 | -91.7218
1.6825 | -92.0782
1.6861 | -92.3717
1.6898 | -92.431
1.6935 | -92.2632
1.6971 | -92.0742
1.7008 | -92.0264
1.7045 | -92.0901
1.7081 | -92.147
1.7118 | -92.1403
1.7155 | -92.1689
1.7191 | -92.321
1.7228 | -92.5708
1.7265 | -92.8379
1.7301 | -93.0879
1.7338 | -93.3291
1.7375 | -93.5449
1.7411 | -93.6552
1.7448 | -93.5322
1.7485 | -93.178
1.7521 | -92.7652
1.7558 | -92.4609
1.7594 | -92.3559
1.7631 | -92.4561
1.7668 | -92.6261
1.7704 | -92.7084
1.7741 | -92.616
1.7778 | -92.4305
1.7814 | -92.2812
1.7851 | -92.2462
1.7888 | -92.3016
1.7924 | -92.3726
1.7961 | -92.4338
1.7998 | -92.5204
1.8034 | -92.6367
1.8071 | -92.7556
1.8108 | -92.8015
1.8144 | -92.7689
1.8181 | -92.7115
1.8218 | -92.65
1.8254 | -92.5674
1.8291 | -92.4445
1.8328 | -92.2991
1.8364 | -92.172
1.8401 | -92.0981
1.8438 | -92.0983
1.8474 | -92.1754
1.8511 | -92.3504
1.8548 | -92.6277
1.8584 | -92.9154
1.8621 | -93.0896
1.8658 | -93.0962
1.8694 | -92.9767
1.8731 | -92.8146
1.8767 | -92.6789
1.8804 | -92.61
1.8841 | -92.6049
1.8877 | -92.6309
1.8914 | -92.6437
1.8951 | -92.6392
1.8987 | -92.6577
1.9024 | -92.7633
1.9061 | -92.9914
1.9097 | -93.2931
1.9134 | -93.5523
1.9171 | -93.6224
1.9207 | -93.4327
1.9244 | -93.1186
1.9281 | -92.8641
1.9317 | -92.7631
1.9354 | -92.7746
1.9391 | -92.7927
1.9427 | -92.7936
1.9464 | -92.7921
1.9501 | -92.7925
1.9537 | -92.7932
1.9574 | -92.802
1.9611 | -92.8338
1.9647 | -92.8587
1.9684 | -92.8762
1.9721 | -92.9611
1.9757 | -93.1347
1.9794 | -93.3378
1.983 | -93.5278
1.9867 | -93.7042
1.9904 | -93.845
1.994 | -93.8795
1.9977 | -93.795
2.0014 | -93.6801
2.005 | -93.6362
2.0087 | -93.6975
2.0124 | -93.8001
2.016 | -93.8661
2.0197 | -93.8931
2.0234 | -93.9039
2.027 | -93.9228
2.0307 | -93.9405
2.0344 | -93.9252
2.038 | -93.8743
2.0417 | -93.8549
2.0454 | -93.9675
2.049 | -94.1976
2.0527 | -94.4405
2.0564 | -94.5681
2.06 | -94.5055
2.0637 | -94.3057
2.0674 | -94.1271
2.071 | -94.089
2.0747 | -94.1664
2.0784 | -94.2136
2.082 | -94.1519
2.0857 | -93.9698
2.0893 | -93.7528
2.093 | -93.6231
2.0967 | -93.6498
2.1003 | -93.7788
2.104 | -93.8908
2.1077 | -93.9094
2.1113 | -93.8981
2.115 | -93.972
2.1187 | -94.0424
2.1223 | -94.14
2.126 | -94.2488
2.1297 | -94.3656
2.1333 | -94.4581
2.137 | -94.4664
2.1407 | -94.3532
2.1443 | -94.1776
2.148 | -94.0147
2.1517 | -93.9245
2.1553 | -93.9176
2.159 | -93.9698
2.1627 | -94.0571
2.1663 | -94.1435
2.17 | -94.2008
2.1737 | -94.2364
2.1773 | -94.298
2.181 | -94.4382
2.1847 | -94.6373
2.1883 | -94.8273
2.192 | -94.9118
2.1956 | -94.8195
2.1993 | -94.5501
2.203 | -94.2023
2.2066 | -93.9383
2.2103 | -93.8528
2.214 | -93.9614
2.2176 | -94.1736
2.2213 | -94.3387
2.225 | -94.3727
2.2286 | -94.2911
2.2323 | -94.1881
2.236 | -94.13
2.2396 | -94.1238
2.2433 | -94.1658
2.247 | -94.2525
2.2506 | -94.3793
2.2543 | -94.4998
2.258 | -94.5837
2.2616 | -94.6352
2.2653 | -94.6823
2.269 | -94.7716
2.2726 | -94.9045
2.2763 | -95.0192
2.28 | -95.0305
2.2836 | -94.8872
2.2873 | -94.641
2.291 | -94.4111
2.2946 | -94.2995
2.2983 | -94.3398
2.3019 | -94.5347
2.3056 | -94.8387
2.3093 | -95.1233
2.3129 | -95.2647
2.3166 | -95.245
2.3203 | -95.1535
2.3239 | -95.0767
2.3276 | -95.0393
2.3313 | -95.0295
2.3349 | -95.0276
2.3386 | -95.0135
2.3423 | -94.9491
2.3459 | -94.8367
2.3496 | -94.7443
2.3533 | -94.7121
2.3569 | -94.7291
2.3606 | -94.7473
2.3643 | -94.7548
2.3679 | -94.7453
2.3716 | -94.7199
2.3753 | -94.6986
2.3789 | -94.7045
2.3826 | -94.7777
2.3863 | -94.9353
2.3899 | -95.1128
2.3936 | -95.241
2.3973 | -95.2666
2.4009 | -95.1862
2.4046 | -95.0403
2.4082 | -94.8874
2.4119 | -94.7469
2.4156 | -94.653
2.4192 | -94.6467
2.4229 | -94.7175
2.4266 | -94.8352
2.4302 | -94.9604
2.4339 | -95.0413
2.4376 | -95.0896
2.4412 | -95.101
2.4449 | -95.0468
2.4486 | -94.9036
2.4522 | -94.6695
2.4559 | -94.3722
2.4596 | -94.0919
2.4632 | -93.9154
2.4669 | -93.8915
2.4706 | -94.0308
2.4742 | -94.3156
2.4779 | -94.6843
2.4816 | -95.1041
2.4852 | -95.5135
2.4889 | -95.8409
2.4926 | -96.089
2.4962 | -96.2829
2.4999 | -96.4116
2.5036 | -96.4376
2.5072 | -96.296
2.5109 | -96.0144
2.5145 | -95.6929
2.5182 | -95.442
2.5219 | -95.3093
2.5255 | -95.3068
2.5292 | -95.3922
2.5329 | -95.4824
2.5365 | -95.5218
2.5402 | -95.5104
2.5439 | -95.4741
2.5475 | -95.4712
2.5512 | -95.5138
2.5549 | -95.5996
2.5585 | -95.6911
2.5622 | -95.7479
2.5659 | -95.76
2.5695 | -95.7177
2.5732 | -95.6673
2.5769 | -95.6713
2.5805 | -95.7703
2.5842 | -95.9461
2.5879 | -96.1293
2.5915 | -96.248
2.5952 | -96.2644
2.5989 | -96.2222
2.6025 | -96.1452
2.6062 | -96.0678
2.6099 | -96.0083
2.6135 | -95.9547
2.6172 | -95.9167
2.6208 | -95.9055
2.6245 | -95.9523
2.6282 | -96.0603
2.6318 | -96.1911
2.6355 | -96.2872
2.6392 | -96.3031
2.6428 | -96.2272
2.6465 | -96.1741
2.6502 | -96.1058
2.6538 | -96.0195
2.6575 | -95.9188
2.6612 | -95.8119
2.6648 | -95.7259
2.6685 | -95.6605
2.6722 | -95.6293
2.6758 | -95.6455
2.6795 | -95.701
2.6832 | -95.7962
2.6868 | -95.9214
2.6905 | -96.072
2.6942 | -96.1964
2.6978 | -96.2574
2.7015 | -96.2427
2.7052 | -96.1474
2.7088 | -96.0061
2.7125 | -95.8601
2.7162 | -95.7403
2.7198 | -95.665
2.7235 | -95.6327
2.7271 | -95.6602
2.7308 | -95.7373
2.7345 | -95.8268
2.7381 | -95.8988
2.7418 | -95.9535
2.7455 | -95.9651
2.7491 | -95.9274
2.7528 | -95.8378
2.7565 | -95.6844
2.7601 | -95.4849
2.7638 | -95.2904
2.7675 | -95.1552
2.7711 | -95.1245
2.7748 | -95.207
2.7785 | -95.4044
2.7821 | -95.6914
2.7858 | -96.0089
2.7895 | -96.2931
2.7931 | -96.5118
2.7968 | -96.6282
2.8005 | -96.6531
2.8041 | -96.647
2.8078 | -96.6431
2.8115 | -96.6069
2.8151 | -96.5319
2.8188 | -96.4043
2.8225 | -96.224
2.8261 | -96.0158
2.8298 | -95.8323
2.8334 | -95.7134
2.8371 | -95.6755
2.8408 | -95.7177
2.8444 | -95.8259
2.8481 | -95.9704
2.8518 | -96.1116
2.8554 | -96.2244
2.8591 | -96.3209
2.8628 | -96.3922
2.8664 | -96.4621
2.8701 | -96.5496
2.8738 | -96.6523
2.8774 | -96.7362
2.8811 | -96.7736
2.8848 | -96.748
2.8884 | -96.6678
2.8921 | -96.5633
2.8958 | -96.4638
2.8994 | -96.3929
2.9031 | -96.3589
2.9068 | -96.3501
2.9104 | -96.3472
2.9141 | -96.3554
2.9178 | -96.3599
2.9214 | -96.3741
2.9251 | -96.4334
2.9288 | -96.5478
2.9324 | -96.7078
2.9361 | -96.8561
2.9397 | -96.9442
2.9434 | -96.95
2.9471 | -96.8677
2.9507 | -96.7192
2.9544 | -96.5503
2.9581 | -96.4065
2.9617 | -96.3066
2.9654 | -96.2826
2.9691 | -96.3198
2.9727 | -96.3938
2.9764 | -96.4694
2.9801 | -96.5416
2.9837 | -96.6086
2.9874 | -96.6609
2.9911 | -96.6915
2.9947 | -96.7126
2.9984 | -96.7314
3.0021 | -96.7488
3.0057 | -96.7583
3.0094 | -96.7592
3.0131 | -96.7536
3.0167 | -96.7792
3.0204 | -96.8526
3.0241 | -96.9923
3.0277 | -97.1502
3.0314 | -97.2585
3.0351 | -97.281
3.0387 | -97.2244
3.0424 | -97.1161
3.046 | -96.9987
3.0497 | -96.9025
3.0534 | -96.8383
3.057 | -96.8005
3.0607 | -96.7723
3.0644 | -96.7379
3.068 | -96.6998
3.0717 | -96.6612
3.0754 | -96.6515
3.079 | -96.6921
3.0827 | -96.7807
3.0864 | -96.915
3.09 | -97.0716
3.0937 | -97.2028
3.0974 | -97.2848
3.101 | -97.3038
3.1047 | -97.2848
3.1084 | -97.268
3.112 | -97.2714
3.1157 | -97.3049
3.1194 | -97.349
3.123 | -97.3626
3.1267 | -97.3069
3.1304 | -97.1894
3.134 | -97.0409
3.1377 | -96.9103
3.1414 | -96.8352
3.145 | -96.8314
3.1487 | -96.888
3.1523 | -96.967
3.156 | -97.0278
3.1597 | -97.05
3.1633 | -97.0116
3.167 | -96.9275
3.1707 | -96.8212
3.1743 | -96.7228
3.178 | -96.6348
3.1817 | -96.5506
3.1853 | -96.486
3.189 | -96.4571
3.1927 | -96.482
3.1963 | -96.556
3.2 | -96.6779
3.2037 | -96.835
3.2073 | -97.0171
3.211 | -97.1808
3.2147 | -97.302
3.2183 | -97.3634
3.222 | -97.3611
3.2257 | -97.3392
3.2293 | -97.3301
3.233 | -97.3411
3.2367 | -97.3775
3.2403 | -97.4161
3.244 | -97.4354
3.2477 | -97.4286
3.2513 | -97.3955
3.255 | -97.3478
3.2586 | -97.2958
3.2623 | -97.2287
3.266 | -97.1623
3.2696 | -97.0972
3.2733 | -97.0354
3.277 | -96.9857
3.2806 | -96.9626
3.2843 | -96.9825
3.288 | -97.0443
3.2916 | -97.1321
3.2953 | -97.2328
3.299 | -97.3172
3.3026 | -97.3789
3.3063 | -97.4266
3.31 | -97.453
3.3136 | -97.4439
3.3173 | -97.4166
3.321 | -97.3883
3.3246 | -97.3725
3.3283 | -97.3838
3.332 | -97.4266
3.3356 | -97.4826
3.3393 | -97.5405
3.343 | -97.5847
3.3466 | -97.5957
3.3503 | -97.5721
3.354 | -97.5311
3.3576 | -97.4736
3.3613 | -97.424
3.3649 | -97.3887
3.3686 | -97.377
3.3723 | -97.3803
3.3759 | -97.3922
3.3796 | -97.4076
3.3833 | -97.4273
3.3869 | -97.4468
3.3906 | -97.4604
3.3943 | -97.4658
3.3979 | -97.4771
3.4016 | -97.4924
3.4053 | -97.4989
3.4089 | -97.5154
3.4126 | -97.5448
3.4163 | -97.5786
3.4199 | -97.6044
3.4236 | -97.6133
3.4273 | -97.6079
3.4309 | -97.5969
3.4346 | -97.6044
3.4383 | -97.6386
3.4419 | -97.7085
3.4456 | -97.8109
3.4493 | -97.9335
3.4529 | -98.0533
3.4566 | -98.1346
3.4603 | -98.1623
3.4639 | -98.113
3.4676 | -97.9928
3.4712 | -97.81
3.4749 | -97.6112
3.4786 | -97.4231
3.4822 | -97.2674
3.4859 | -97.1629
3.4896 | -97.1124
3.4932 | -97.1141
3.4969 | -97.152
3.5006 | -97.2228
3.5042 | -97.3042
3.5079 | -97.382
3.5116 | -97.4674
3.5152 | -97.5436
3.5189 | -97.6054
3.5226 | -97.6781
3.5262 | -97.7651
3.5299 | -97.85
3.5336 | -97.9212
3.5372 | -97.9627
3.5409 | -97.9737
3.5446 | -97.9461
3.5482 | -97.89
3.5519 | -97.8189
3.5556 | -97.7465
3.5592 | -97.6917
3.5629 | -97.6622
3.5666 | -97.647
3.5702 | -97.6585
3.5739 | -97.6988
3.5775 | -97.7712
3.5812 | -97.8841
3.5849 | -98.0304
3.5885 | -98.2032
3.5922 | -98.3944
3.5959 | -98.5765
3.5995 | -98.7425
3.6032 | -98.8756
3.6069 | -98.9642
3.6105 | -98.9894
3.6142 | -98.9423
3.6179 | -98.8564
3.6215 | -98.7637
3.6252 | -98.6699
3.6289 | -98.5825
3.6325 | -98.4937
3.6362 | -98.4148
3.6399 | -98.3544
3.6435 | -98.3221
3.6472 | -98.318
3.6509 | -98.3516
3.6545 | -98.4105
3.6582 | -98.4779
3.6619 | -98.5287
3.6655 | -98.5417
3.6692 | -98.5192
3.6729 | -98.4607
3.6765 | -98.406
3.6802 | -98.3856
3.6838 | -98.4161
3.6875 | -98.4988
3.6912 | -98.5929
3.6948 | -98.6688
3.6985 | -98.6567
3.7022 | -98.6668
3.7058 | -98.6627
3.7095 | -98.6531
3.7132 | -98.6462
3.7168 | -98.6471
3.7205 | -98.6578
3.7242 | -98.6719
3.7278 | -98.674
3.7315 | -98.6459
3.7352 | -98.5869
3.7388 | -98.5149
3.7425 | -98.4323
3.7462 | -98.3423
3.7498 | -98.2621
3.7535 | -98.182
3.7572 | -98.1034
3.7608 | -98.0198
3.7645 | -97.9373
3.7682 | -97.8622
3.7718 | -97.8071
3.7755 | -97.7736
3.7792 | -97.7587
3.7828 | -97.7675
3.7865 | -97.7885
3.7901 | -97.8146
3.7938 | -97.8487
3.7975 | -97.8834
3.8011 | -97.9014
3.8048 | -97.9023
3.8085 | -97.8955
3.8121 | -97.8972
3.8158 | -97.9129
3.8195 | -97.9427
3.8231 | -97.9971
3.8268 | -98.0813
3.8305 | -98.191
3.8341 | -98.3107
3.8378 | -98.4389
3.8415 | -98.5661
3.8451 | -98.6828
3.8488 | -98.7945
3.8525 | -98.8822
3.8561 | -98.9405
3.8598 | -98.9682
3.8635 | -98.9684
3.8671 | -98.936
3.8708 | -98.8818
3.8745 | -98.8002
3.8781 | -98.7127
3.8818 | -98.631
3.8855 | -98.5587
3.8891 | -98.4886
3.8928 | -98.4232
3.8964 | -98.3772
3.9001 | -98.3498
3.9038 | -98.3451
3.9074 | -98.3654
3.9111 | -98.4172
3.9148 | -98.5091
3.9184 | -98.6317
3.9221 | -98.7726
3.9258 | -98.9193
3.9294 | -99.0485
3.9331 | -99.1286
3.9368 | -99.1627
3.9404 | -99.1353
3.9441 | -99.0536
3.9478 | -98.9498
3.9514 | -98.8496
3.9551 | -98.774
3.9588 | -98.7324
3.9624 | -98.7232
3.9661 | -98.738
3.9698 | -98.7716
3.9734 | -98.8117
3.9771 | -98.8551
3.9808 | -98.9046
3.9844 | -98.9642
3.9881 | -99.0307
3.9918 | -99.0909
3.9954 | -99.1367
3.9991 | -99.1612
4.0027 | -99.1768
4.0064 | -99.1867
4.0101 | -99.1853
4.0137 | -99.1902
4.0174 | -99.2059
4.0211 | -99.2343
4.0247 | -99.2771
4.0284 | -99.3292
4.0321 | -99.384
4.0357 | -99.4285
4.0394 | -99.4596
4.0431 | -99.477
4.0467 | -99.4772
4.0504 | -99.4621
4.0541 | -99.4312
4.0577 | -99.3857
4.0614 | -99.343
4.0651 | -99.2994
4.0687 | -99.2662
4.0724 | -99.2526
4.0761 | -99.2495
4.0797 | -99.2491
4.0834 | -99.2455
4.0871 | -99.2367
4.0907 | -99.2094
4.0944 | -99.1661
4.0981 | -99.1248
4.1017 | -99.0939
4.1054 | -99.0775
4.109 | -99.0912
4.1127 | -99.1361
4.1164 | -99.2153
4.12 | -99.3124
4.1237 | -99.4038
4.1274 | -99.4839
4.131 | -99.5353
4.1347 | -99.5515
4.1384 | -99.5415
4.142 | -99.516
4.1457 | -99.4875
4.1494 | -99.4652
4.153 | -99.4505
4.1567 | -99.4331
4.1604 | -99.4093
4.164 | -99.3719
4.1677 | -99.3197
4.1714 | -99.2634
4.175 | -99.2141
4.1787 | -99.1784
4.1824 | -99.1596
4.186 | -99.1492
4.1897 | -99.1422
4.1934 | -99.1314
4.197 | -99.1237
4.2007 | -99.1222
4.2044 | -99.1303
4.208 | -99.1444
4.2117 | -99.1605
4.2153 | -99.1752
4.219 | -99.1949
4.2227 | -99.2186
4.2263 | -99.2536
4.23 | -99.2906
4.2337 | -99.3305
4.2373 | -99.3657
4.241 | -99.3987
4.2447 | -99.4268
4.2483 | -99.4655
4.252 | -99.5148
4.2557 | -99.5666
4.2593 | -99.6283
4.263 | -99.6866
4.2667 | -99.7246
4.2703 | -99.7457
4.274 | -99.7433
4.2777 | -99.7086
4.2813 | -99.6604
4.285 | -99.5963
4.2887 | -99.5266
4.2923 | -99.4728
4.296 | -99.4443
4.2997 | -99.4408
4.3033 | -99.4699
4.307 | -99.5272
4.3107 | -99.5977
4.3143 | -99.6814
4.318 | -99.7605
4.3216 | -99.8392
4.3253 | -99.914
4.329 | -99.9701
4.3326 | -100.0042
4.3363 | -100.0176
4.34 | -100.0106
4.3436 | -99.9855
4.3473 | -99.9468
4.351 | -99.8845
4.3546 | -99.8035
4.3583 | -99.7289
4.362 | -99.6591
4.3656 | -99.5962
4.3693 | -99.5427
4.373 | -99.4889
4.3766 | -99.4402
4.3803 | -99.4028
4.384 | -99.3789
4.3876 | -99.3797
4.3913 | -99.4084
4.395 | -99.4601
4.3986 | -99.5301
4.4023 | -99.6055
4.406 | -99.6709
4.4096 | -99.7097
4.4133 | -99.7225
4.417 | -99.7119
4.4206 | -99.679
4.4243 | -99.6287
4.428 | -99.5495
4.4316 | -99.455
4.4353 | -99.3537
4.4389 | -99.2552
4.4426 | -99.1754
4.4463 | -99.1161
4.4499 | -99.0936
4.4536 | -99.1087
4.4573 | -99.1592
4.4609 | -99.237
4.4646 | -99.3347
4.4683 | -99.4427
4.4719 | -99.549
4.4756 | -99.6389
4.4793 | -99.7004
4.4829 | -99.7313
4.4866 | -99.7355
4.4903 | -99.7135
4.4939 | -99.6666
4.4976 | -99.6115
4.5013 | -99.5525
4.5049 | -99.4958
4.5086 | -99.4522
4.5123 | -99.4141
4.5159 | -99.3901
4.5196 | -99.3846
4.5233 | -99.4009
4.5269 | -99.4402
4.5306 | -99.5115
4.5343 | -99.6074
4.5379 | -99.7189
4.5416 | -99.8421
4.5452 | -99.9759
4.5489 | -100.1009
4.5526 | -100.2089
4.5562 | -100.2877
4.5599 | -100.3328
4.5636 | -100.3529
4.5672 | -100.33
4.5709 | -100.2852
4.5746 | -100.2297
4.5782 | -100.1579
4.5819 | -100.0867
4.5856 | -100.0176
4.5892 | -99.9438
4.5929 | -99.866
4.5966 | -99.7922
4.6002 | -99.722
4.6039 | -99.6634
4.6076 | -99.6177
4.6112 | -99.591
4.6149 | -99.585
4.6186 | -99.6113
4.6222 | -99.6633
4.6259 | -99.7289
4.6296 | -99.8091
4.6332 | -99.8839
4.6369 | -99.947
4.6406 | -99.9905
4.6442 | -100.012
4.6479 | -100.031
4.6515 | -100.0579
4.6552 | -100.0899
4.6589 | -100.1331
4.6625 | -100.1961
4.6662 | -100.2606
4.6699 | -100.3176
4.6735 | -100.3757
4.6772 | -100.4288
4.6809 | -100.4652
4.6845 | -100.4815
4.6882 | -100.4791
4.6919 | -100.4631
4.6955 | -100.4476
4.6992 | -100.4278
4.7029 | -100.4142
4.7065 | -100.4101
4.7102 | -100.4141
4.7139 | -100.431
4.7175 | -100.4499
4.7212 | -100.4771
4.7249 | -100.5151
4.7285 | -100.5605
4.7322 | -100.6042
4.7359 | -100.6513
4.7395 | -100.6923
4.7432 | -100.7126
4.7469 | -100.7185
4.7505 | -100.7167
4.7542 | -100.7093
4.7578 | -100.7132
4.7615 | -100.7164
4.7652 | -100.7213
4.7688 | -100.7239
4.7725 | -100.7257
4.7762 | -100.7258
4.7798 | -100.7179
4.7835 | -100.6963
4.7872 | -100.6587
4.7908 | -100.6184
4.7945 | -100.5649
4.7982 | -100.5184
4.8018 | -100.4782
4.8055 | -100.446
4.8092 | -100.4191
4.8128 | -100.3988
4.8165 | -100.3865
4.8202 | -100.3808
4.8238 | -100.3764
4.8275 | -100.3695
4.8312 | -100.3591
4.8348 | -100.3563
4.8385 | -100.3628
4.8422 | -100.3749
4.8458 | -100.3902
4.8495 | -100.4112
4.8532 | -100.4395
4.8568 | -100.4727
4.8605 | -100.5163
4.8641 | -100.5638
4.8678 | -100.6131
4.8715 | -100.6734
4.8751 | -100.7235
4.8788 | -100.7662
4.8825 | -100.8034
4.8861 | -100.8234
4.8898 | -100.8279
4.8935 | -100.8105
4.8971 | -100.7748
4.9008 | -100.7227
4.9045 | -100.6741
4.9081 | -100.622
4.9118 | -100.5686
4.9155 | -100.5214
4.9191 | -100.4898
4.9228 | -100.4711
4.9265 | -100.4609
4.9301 | -100.4636
4.9338 | -100.4804
4.9375 | -100.506
4.9411 | -100.5457
4.9448 | -100.5808
4.9485 | -100.6138
4.9521 | -100.6543
4.9558 | -100.6941
4.9595 | -100.7303
4.9631 | -100.7612
4.9668 | -100.7857
4.9704 | -100.8097
4.9741 | -100.8282
4.9778 | -100.8512
4.9814 | -100.8622
4.9851 | -100.876
4.9888 | -100.8843
4.9924 | -100.8879
4.9961 | -100.8765
4.9998 | -100.86
5.0034 | -100.8282
5.0071 | -100.7846
5.0108 | -100.7348
5.0144 | -100.6803
5.0181 | -100.632
5.0218 | -100.6022
5.0254 | -100.5822
5.0291 | -100.5582
5.0328 | -100.5406
5.0364 | -100.5278
5.0401 | -100.5258
5.0438 | -100.529
5.0474 | -100.5341
5.0511 | -100.5461
5.0548 | -100.5614
5.0584 | -100.5897
5.0621 | -100.6286
5.0658 | -100.6744
5.0694 | -100.718
5.0731 | -100.7636
5.0767 | -100.7995
5.0804 | -100.8328
5.0841 | -100.8544
5.0877 | -100.8702
5.0914 | -100.8749
5.0951 | -100.868
5.0987 | -100.8535
5.1024 | -100.8401
5.1061 | -100.8346
5.1097 | -100.8299
5.1134 | -100.8251
5.1171 | -100.8271
5.1207 | -100.8368
5.1244 | -100.8543
5.1281 | -100.8824
5.1317 | -100.92
5.1354 | -100.965
5.1391 | -101.0178
5.1427 | -101.0764
5.1464 | -101.1361
5.1501 | -101.1923
5.1537 | -101.2462
5.1574 | -101.2985
5.1611 | -101.3488
5.1647 | -101.3907
5.1684 | -101.4364
5.172 | -101.4831
5.1757 | -101.5345
5.1794 | -101.5839
5.183 | -101.6336
5.1867 | -101.6869
5.1904 | -101.7419
5.194 | -101.7791
5.1977 | -101.7997
5.2014 | -101.7994
5.205 | -101.7835
5.2087 | -101.7483
5.2124 | -101.6871
5.216 | -101.6101
5.2197 | -101.5284
5.2234 | -101.4496
5.227 | -101.3767
5.2307 | -101.3098
5.2344 | -101.2578
5.238 | -101.2246
5.2417 | -101.2068
5.2454 | -101.191
5.249 | -101.1801
5.2527 | -101.1659
5.2564 | -101.1536
5.26 | -101.1411
5.2637 | -101.1313
5.2674 | -101.1205
5.271 | -101.1175
5.2747 | -101.1228
5.2784 | -101.1239
5.282 | -101.1161
5.2857 | -101.0936
5.2893 | -101.073
5.293 | -101.0496
5.2967 | -101.0273
5.3003 | -101.004
5.304 | -100.992
5.3077 | -100.9915
5.3113 | -101.0053
5.315 | -101.023
5.3187 | -101.0471
5.3223 | -101.0756
5.326 | -101.1079
5.3297 | -101.1345
5.3333 | -101.1562
5.337 | -101.1719
5.3407 | -101.1822
5.3443 | -101.1848
5.348 | -101.1815
5.3517 | -101.1726
5.3553 | -101.1486
5.359 | -101.1198
5.3627 | -101.0869
5.3663 | -101.0601
5.37 | -101.0371
5.3737 | -101.0227
5.3773 | -101.0215
5.381 | -101.0348
5.3847 | -101.0626
5.3883 | -101.1008
5.392 | -101.1534
5.3956 | -101.2213
5.3993 | -101.2965
5.403 | -101.3717
5.4066 | -101.4409
5.4103 | -101.5044
5.414 | -101.5606
5.4176 | -101.607
5.4213 | -101.6309
5.425 | -101.6377
5.4286 | -101.6279
5.4323 | -101.5993
5.436 | -101.5521
5.4396 | -101.4968
5.4433 | -101.4387
5.447 | -101.3747
5.4506 | -101.3138
5.4543 | -101.2661
5.458 | -101.2342
5.4616 | -101.2147
5.4653 | -101.2029
5.469 | -101.2063
5.4726 | -101.2217
5.4763 | -101.2453
5.48 | -101.2765
5.4836 | -101.3175
5.4873 | -101.3639
5.491 | -101.4181
5.4946 | -101.4657
5.4983 | -101.504
5.5019 | -101.5421
5.5056 | -101.5752
5.5093 | -101.5982
5.5129 | -101.6162
5.5166 | -101.6413
5.5203 | -101.6684
5.5239 | -101.698
5.5276 | -101.7369
5.5313 | -101.7894
5.5349 | -101.845
5.5386 | -101.8989
5.5423 | -101.9504
5.5459 | -101.9997
5.5496 | -102.0437
5.5533 | -102.0677
5.5569 | -102.0818
5.5606 | -102.0831
5.5643 | -102.0687
5.5679 | -102.0426
5.5716 | -102.0091
5.5753 | -101.9711
5.5789 | -101.937
5.5826 | -101.9103
5.5863 | -101.8834
5.5899 | -101.8667
5.5936 | -101.8505
5.5973 | -101.8353
5.6009 | -101.8173
5.6046 | -101.8
5.6082 | -101.7892
5.6119 | -101.7722
5.6156 | -101.7597
5.6192 | -101.7614
5.6229 | -101.7741
5.6266 | -101.8007
5.6302 | -101.8383
5.6339 | -101.8856
5.6376 | -101.9408
5.6412 | -101.988
5.6449 | -102.0354
5.6486 | -102.0804
5.6522 | -102.1066
5.6559 | -102.117
5.6596 | -102.1125
5.6632 | -102.0926
5.6669 | -102.0665
5.6706 | -102.0333
5.6742 | -101.9942
5.6779 | -101.9595
5.6816 | -101.9267
5.6852 | -101.9009
5.6889 | -101.8824
5.6926 | -101.8674
5.6962 | -101.866
5.6999 | -101.8701
5.7036 | -101.8766
5.7072 | -101.8918
5.7109 | -101.9042
5.7145 | -101.9173
5.7182 | -101.9338
5.7219 | -101.9502
5.7255 | -101.9597
5.7292 | -101.9506
5.7329 | -101.9216
5.7365 | -101.8916
5.7402 | -101.8504
5.7439 | -101.8073
5.7475 | -101.7656
5.7512 | -101.719
5.7549 | -101.6719
5.7585 | -101.6345
5.7622 | -101.6078
5.7659 | -101.5868
5.7695 | -101.5651
5.7732 | -101.5454
5.7769 | -101.5276
5.7805 | -101.5094
5.7842 | -101.4946
5.7879 | -101.4807
5.7915 | -101.4671
5.7952 | -101.466
5.7989 | -101.4765
5.8025 | -101.4985
5.8062 | -101.5359
5.8099 | -101.5845
5.8135 | -101.6558
5.8172 | -101.7113
5.8208 | -101.7706
5.8245 | -101.8275
5.8282 | -101.8831
5.8318 | -101.9274
5.8355 | -101.9686
5.8392 | -102.0072
5.8428 | -102.0404
5.8465 | -102.0569
5.8502 | -102.0554
5.8538 | -102.0468
5.8575 | -102.0312
5.8612 | -102.0114
5.8648 | -101.9973
5.8685 | -101.9679
5.8722 | -101.9226
5.8758 | -101.8777
5.8795 | -101.8316
5.8832 | -101.7701
5.8868 | -101.7006
5.8905 | -101.624
5.8942 | -101.5515
5.8978 | -101.4834
5.9015 | -101.4191
5.9052 | -101.3704
5.9088 | -101.338
5.9125 | -101.3199
5.9162 | -101.3097
5.9198 | -101.3085
5.9235 | -101.3251
5.9271 | -101.3502
5.9308 | -101.3796
5.9345 | -101.4145
5.9381 | -101.4569
5.9418 | -101.502
5.9455 | -101.5465
5.9491 | -101.5929
5.9528 | -101.6366
5.9565 | -101.67
5.9601 | -101.7135
5.9638 | -101.756
5.9675 | -101.7804
5.9711 | -101.8011
5.9748 | -101.8284
5.9785 | -101.8484
5.9821 | -101.8654
5.9858 | -101.8721
5.9895 | -101.8761
5.9931 | -101.8789
5.9968 | -101.8692
6.0005 | -101.8487
6.0041 | -101.8213
6.0078 | -101.7963
6.0115 | -101.7744
6.0151 | -101.7529
6.0188 | -101.7335
6.0225 | -101.7155
6.0261 | -101.6922
6.0298 | -101.6664
6.0334 | -101.6448
6.0371 | -101.622
6.0408 | -101.6025
6.0444 | -101.5915
6.0481 | -101.5918
6.0518 | -101.6011
6.0554 | -101.6236
6.0591 | -101.6615
6.0628 | -101.6972
6.0664 | -101.7272
6.0701 | -101.765
6.0738 | -101.7921
6.0774 | -101.8079
6.0811 | -101.8138
6.0848 | -101.8081
6.0884 | -101.7997
6.0921 | -101.7914
6.0958 | -101.7791
6.0994 | -101.7568
6.1031 | -101.7289
6.1068 | -101.6965
6.1104 | -101.6699
6.1141 | -101.6532
6.1178 | -101.6497
6.1214 | -101.6531
6.1251 | -101.66
6.1288 | -101.6769
6.1324 | -101.7032
6.1361 | -101.7322
6.1397 | -101.7667
6.1434 | -101.8018
6.1471 | -101.8337
6.1507 | -101.8607
6.1544 | -101.8835
6.1581 | -101.9007
6.1617 | -101.9129
6.1654 | -101.9379
6.1691 | -101.971
6.1727 | -102.0069
6.1764 | -102.0429
6.1801 | -102.0752
6.1837 | -102.1029
6.1874 | -102.1238
6.1911 | -102.142
6.1947 | -102.1506
6.1984 | -102.145
6.2021 | -102.1271
6.2057 | -102.1036
6.2094 | -102.0793
6.2131 | -102.0593
6.2167 | -102.0456
6.2204 | -102.0403
6.2241 | -102.0398
6.2277 | -102.0441
6.2314 | -102.0515
6.2351 | -102.0658
6.2387 | -102.0876
6.2424 | -102.11
6.246 | -102.1352
6.2497 | -102.1617
6.2534 | -102.1755
6.257 | -102.1773
6.2607 | -102.171
6.2644 | -102.1583
6.268 | -102.147
6.2717 | -102.1274
6.2754 | -102.104
6.279 | -102.0816
6.2827 | -102.0562
6.2864 | -102.032
6.29 | -102.0172
6.2937 | -102.009
6.2974 | -102.0083
6.301 | -102.0192
6.3047 | -102.0441
6.3084 | -102.0847
6.312 | -102.1347
6.3157 | -102.1956
6.3194 | -102.2629
6.323 | -102.3321
6.3267 | -102.3975
6.3304 | -102.4546
6.334 | -102.5063
6.3377 | -102.5535
6.3414 | -102.5889
6.345 | -102.611
6.3487 | -102.6453
6.3523 | -102.6793
6.356 | -102.7154
6.3597 | -102.7468
6.3633 | -102.7748
6.367 | -102.8048
6.3707 | -102.8266
6.3743 | -102.8482
6.378 | -102.8694
6.3817 | -102.8827
6.3853 | -102.8826
6.389 | -102.876
6.3927 | -102.8635
6.3963 | -102.8436
6.4 | -102.8131
6.4037 | -102.7702
6.4073 | -102.7318
6.411 | -102.6884
6.4147 | -102.6484
6.4183 | -102.6119
6.422 | -102.5757
6.4257 | -102.5434
6.4293 | -102.5135
6.433 | -102.4903
6.4367 | -102.4785
6.4403 | -102.47
6.444 | -102.4703
6.4477 | -102.4802
6.4513 | -102.4951
6.455 | -102.5156
6.4586 | -102.5407
6.4623 | -102.5647
6.466 | -102.5844
6.4696 | -102.6015
6.4733 | -102.6085
6.477 | -102.6071
6.4806 | -102.6045
6.4843 | -102.5978
6.488 | -102.5807
6.4916 | -102.5604
6.4953 | -102.5419
6.499 | -102.53
6.5026 | -102.531
6.5063 | -102.5325
6.51 | -102.5458
6.5136 | -102.5696
6.5173 | -102.5956
6.521 | -102.627
6.5246 | -102.6656
6.5283 | -102.7067
6.532 | -102.7404
6.5356 | -102.7677
6.5393 | -102.7892
6.543 | -102.796
6.5466 | -102.7839
6.5503 | -102.7613
6.554 | -102.7351
6.5576 | -102.711
6.5613 | -102.6859
6.5649 | -102.6614
6.5686 | -102.641
6.5723 | -102.6267
6.5759 | -102.6199
6.5796 | -102.6206
6.5833 | -102.6283
6.5869 | -102.6402
6.5906 | -102.6531
6.5943 | -102.6655
6.5979 | -102.6762
6.6016 | -102.6816
6.6053 | -102.6836
6.6089 | -102.6812
6.6126 | -102.6712
6.6163 | -102.6613
6.6199 | -102.6537
6.6236 | -102.6443
6.6273 | -102.6381
6.6309 | -102.6359
6.6346 | -102.6397
6.6383 | -102.6548
6.6419 | -102.6734
6.6456 | -102.7006
6.6493 | -102.7331
6.6529 | -102.7611
6.6566 | -102.7791
6.6603 | -102.7851
6.6639 | -102.7908
6.6676 | -102.7861
6.6712 | -102.77
6.6749 | -102.7446
6.6786 | -102.7103
6.6822 | -102.6702
6.6859 | -102.6269
6.6896 | -102.5771
6.6932 | -102.5252
6.6969 | -102.4774
6.7006 | -102.4401
6.7042 | -102.4072
6.7079 | -102.3817
6.7116 | -102.3706
6.7152 | -102.3662
6.7189 | -102.3661
6.7226 | -102.3748
6.7262 | -102.3857
6.7299 | -102.4038
6.7336 | -102.4239
6.7372 | -102.4353
6.7409 | -102.4428
6.7446 | -102.4517
6.7482 | -102.4541
6.7519 | -102.4525
6.7556 | -102.4519
6.7592 | -102.4487
6.7629 | -102.4363
6.7666 | -102.4239
6.7702 | -102.4139
6.7739 | -102.4051
6.7775 | -102.4017
6.7812 | -102.4032
6.7849 | -102.4148
6.7885 | -102.4349
6.7922 | -102.4601
6.7959 | -102.4934
6.7995 | -102.5213
6.8032 | -102.5521
6.8069 | -102.5839
6.8105 | -102.6086
6.8142 | -102.6218
6.8179 | -102.6186
6.8215 | -102.6061
6.8252 | -102.5766
6.8289 | -102.5306
6.8325 | -102.4839
6.8362 | -102.4355
6.8399 | -102.3862
6.8435 | -102.3319
6.8472 | -102.2875
6.8509 | -102.2525
6.8545 | -102.2318
6.8582 | -102.2262
6.8619 | -102.2351
6.8655 | -102.253
6.8692 | -102.28
6.8729 | -102.3424
6.8765 | -102.3971
6.8802 | -102.4529
6.8838 | -102.5117
6.8875 | -102.5796
6.8912 | -102.6484
6.8948 | -102.7174
6.8985 | -102.7798
6.9022 | -102.8451
6.9058 | -102.9092
6.9095 | -102.9665
6.9132 | -103.013
6.9168 | -103.0574
6.9205 | -103.0981
6.9242 | -103.1302
6.9278 | -103.1669
6.9315 | -103.199
6.9352 | -103.2109
6.9388 | -103.2195
6.9425 | -103.2256
6.9462 | -103.2354
6.9498 | -103.2374
6.9535 | -103.2415
6.9572 | -103.2428
6.9608 | -103.2428
6.9645 | -103.2353
6.9682 | -103.2227
6.9718 | -103.2028
6.9755 | -103.1755
6.9792 | -103.136
6.9828 | -103.0875
6.9865 | -103.0417
6.9901 | -102.9953
6.9938 | -102.9423
6.9975 | -102.8985
7.0011 | -102.8581
7.0048 | -102.8224
7.0085 | -102.7879
7.0121 | -102.761
7.0158 | -102.7454
7.0195 | -102.7372
7.0231 | -102.7341
7.0268 | -102.7361
7.0305 | -102.7416
7.0341 | -102.7555
7.0378 | -102.7744
7.0415 | -102.8052
7.0451 | -102.8356
7.0488 | -102.8684
7.0525 | -102.9004
7.0561 | -102.9387
7.0598 | -102.9731
7.0635 | -103.009
7.0671 | -103.0459
7.0708 | -103.0814
7.0745 | -103.1154
7.0781 | -103.1438
7.0818 | -103.1578
7.0855 | -103.1631
7.0891 | -103.1573
7.0928 | -103.1429
7.0964 | -103.1154
7.1001 | -103.082
7.1038 | -103.0441
7.1074 | -103.001
7.1111 | -102.9515
7.1148 | -102.8989
7.1184 | -102.85
7.1221 | -102.8006
7.1258 | -102.7606
7.1294 | -102.7264
7.1331 | -102.6981
7.1368 | -102.6803
7.1404 | -102.6659
7.1441 | -102.6623
7.1478 | -102.672
7.1514 | -102.7002
7.1551 | -102.7325
7.1588 | -102.769
7.1624 | -102.8108
7.1661 | -102.8598
7.1698 | -102.9072
7.1734 | -102.9516
7.1771 | -102.9965
7.1808 | -103.0388
7.1844 | -103.0766
7.1881 | -103.1104
7.1918 | -103.1386
7.1954 | -103.1653
7.1991 | -103.1854
7.2027 | -103.2069
7.2064 | -103.2249
7.2101 | -103.2414
7.2137 | -103.2613
7.2174 | -103.2843
7.2211 | -103.3036
7.2247 | -103.3144
7.2284 | -103.3289
7.2321 | -103.346
7.2357 | -103.3677
7.2394 | -103.388
7.2431 | -103.4022
7.2467 | -103.4129
7.2504 | -103.418
7.2541 | -103.4209
7.2577 | -103.4129
7.2614 | -103.3977
7.2651 | -103.3733
7.2687 | -103.3412
7.2724 | -103.3029
7.2761 | -103.2628
7.2797 | -103.2235
7.2834 | -103.1801
7.2871 | -103.1333
7.2907 | -103.0938
7.2944 | -103.0636
7.2981 | -103.0395
7.3017 | -103.0176
7.3054 | -103.0008
7.309 | -102.9975
7.3127 | -103.0067
7.3164 | -103.0259
7.32 | -103.0489
7.3237 | -103.0745
7.3274 | -103.1078
7.331 | -103.1369
7.3347 | -103.1619
7.3384 | -103.1785
7.342 | -103.1805
7.3457 | -103.172
7.3494 | -103.1528
7.353 | -103.1266
7.3567 | -103.0886
7.3604 | -103.0379
7.364 | -102.9933
7.3677 | -102.9576
7.3714 | -102.9235
7.375 | -102.8966
7.3787 | -102.8763
7.3824 | -102.8633
7.386 | -102.8624
7.3897 | -102.8682
7.3934 | -102.8841
7.397 | -102.8975
7.4007 | -102.9115
7.4044 | -102.9245
7.408 | -102.9421
7.4117 | -102.9641
7.4153 | -102.9893
7.419 | -103.011
7.4227 | -103.0344
7.4263 | -103.0621
7.43 | -103.0998
7.4337 | -103.1424
7.4373 | -103.1849
7.441 | -103.2318
7.4447 | -103.2815
7.4483 | -103.3294
7.452 | -103.3776
7.4557 | -103.4284
7.4593 | -103.4804
7.463 | -103.5279
7.4667 | -103.5722
7.4703 | -103.6106
7.474 | -103.6365
7.4777 | -103.6524
7.4813 | -103.6626
7.485 | -103.6634
7.4887 | -103.6603
7.4923 | -103.654
7.496 | -103.642
7.4997 | -103.639
7.5033 | -103.6363
7.507 | -103.6373
7.5107 | -103.6365
7.5143 | -103.6337
7.518 | -103.6262
7.5216 | -103.6186
7.5253 | -103.6088
7.529 | -103.596
7.5326 | -103.5774
7.5363 | -103.55
7.54 | -103.5184
7.5436 | -103.4799
7.5473 | -103.4363
7.551 | -103.3913
7.5546 | -103.3473
7.5583 | -103.3026
7.562 | -103.2673
7.5656 | -103.2374
7.5693 | -103.2104
7.573 | -103.1877
7.5766 | -103.1748
7.5803 | -103.1671
7.584 | -103.1598
7.5876 | -103.1566
7.5913 | -103.1598
7.595 | -103.1668
7.5986 | -103.1714
7.6023 | -103.1689
7.606 | -103.1655
7.6096 | -103.1619
7.6133 | -103.1524
7.617 | -103.1381
7.6206 | -103.1251
7.6243 | -103.1206
7.628 | -103.1172
7.6316 | -103.1096
7.6353 | -103.1098
7.6389 | -103.1185
7.6426 | -103.1297
7.6463 | -103.1499
7.6499 | -103.1823
7.6536 | -103.2211
7.6573 | -103.2665
7.6609 | -103.3147
7.6646 | -103.3597
7.6683 | -103.4015
7.6719 | -103.4438
7.6756 | -103.4778
7.6793 | -103.5057
7.6829 | -103.5262
7.6866 | -103.5391
7.6903 | -103.5419
7.6939 | -103.5317
7.6976 | -103.5168
7.7013 | -103.4945
7.7049 | -103.4638
7.7086 | -103.4289
7.7123 | -103.3977
7.7159 | -103.3712
7.7196 | -103.3458
7.7233 | -103.3241
7.7269 | -103.3116
7.7306 | -103.3071
7.7343 | -103.3039
7.7379 | -103.308
7.7416 | -103.3209
7.7452 | -103.3499
7.7489 | -103.3791
7.7526 | -103.4102
7.7562 | -103.4504
7.7599 | -103.4919
7.7636 | -103.5337
7.7672 | -103.5719
7.7709 | -103.6081
7.7746 | -103.6438
7.7782 | -103.6747
7.7819 | -103.6934
7.7856 | -103.7065
7.7892 | -103.7131
7.7929 | -103.7178
7.7966 | -103.7171
7.8002 | -103.7133
7.8039 | -103.7102
7.8076 | -103.7162
7.8112 | -103.723
7.8149 | -103.7291
7.8186 | -103.733
7.8222 | -103.7399
7.8259 | -103.7511
7.8296 | -103.765
7.8332 | -103.7769
7.8369 | -103.7904
7.8406 | -103.8042
7.8442 | -103.8226
7.8479 | -103.8462
7.8515 | -103.8643
7.8552 | -103.8814
7.8589 | -103.8929
7.8625 | -103.9076
7.8662 | -103.9275
7.8699 | -103.9467
7.8735 | -103.9619
7.8772 | -103.9753
7.8809 | -103.9865
7.8845 | -103.9955
7.8882 | -104.0021
7.8919 | -104.0092
7.8955 | -104.0159
7.8992 | -104.0133
7.9029 | -104.0082
7.9065 | -103.9927
7.9102 | -103.9733
7.9139 | -103.9549
7.9175 | -103.9344
7.9212 | -103.9096
7.9249 | -103.881
7.9285 | -103.8508
7.9322 | -103.8277
7.9359 | -103.8113
7.9395 | -103.7903
7.9432 | -103.7698
7.9469 | -103.7632
7.9505 | -103.7637
7.9542 | -103.7686
7.9578 | -103.7795
7.9615 | -103.791
7.9652 | -103.8039
7.9688 | -103.8226
7.9725 | -103.846
7.9762 | -103.8757
7.9798 | -103.9026
7.9835 | -103.9198
7.9872 | -103.9323
7.9908 | -103.9445
7.9945 | -103.9499
7.9982 | -103.9485
8.0018 | -103.9419
8.0055 | -103.9292
8.0092 | -103.9065
8.0128 | -103.882
8.0165 | -103.8546
8.0202 | -103.8222
8.0238 | -103.7906
8.0275 | -103.7596
8.0312 | -103.7222
8.0348 | -103.6878
8.0385 | -103.6527
8.0422 | -103.6188
8.0458 | -103.5943
8.0495 | -103.5766
8.0532 | -103.556
8.0568 | -103.5354
8.0605 | -103.5233
8.0641 | -103.513
8.0678 | -103.4993
8.0715 | -103.4956
8.0751 | -103.4974
8.0788 | -103.4984
8.0825 | -103.509
8.0861 | -103.5253
8.0898 | -103.5395
8.0935 | -103.5553
8.0971 | -103.5713
8.1008 | -103.5917
8.1045 | -103.615
8.1081 | -103.6423
8.1118 | -103.6713
8.1155 | -103.7019
8.1191 | -103.7343
8.1228 | -103.7644
8.1265 | -103.7944
8.1301 | -103.824
8.1338 | -103.8485
8.1375 | -103.8666
8.1411 | -103.885
8.1448 | -103.8968
8.1485 | -103.8984
8.1521 | -103.8989
8.1558 | -103.892
8.1594 | -103.8841
8.1631 | -103.8673
8.1668 | -103.8409
8.1704 | -103.8134
8.1741 | -103.7881
8.1778 | -103.772
8.1814 | -103.7556
8.1851 | -103.7408
8.1888 | -103.7337
8.1924 | -103.7247
8.1961 | -103.7215
8.1998 | -103.7238
8.2034 | -103.725
8.2071 | -103.7272
8.2108 | -103.733
8.2144 | -103.7429
8.2181 | -103.7508
8.2218 | -103.7579
8.2254 | -103.7628
8.2291 | -103.7593
8.2328 | -103.7489
8.2364 | -103.7374
8.2401 | -103.7275
8.2438 | -103.715
8.2474 | -103.7005
8.2511 | -103.6875
8.2548 | -103.6784
8.2584 | -103.6764
8.2621 | -103.675
8.2658 | -103.6748
8.2694 | -103.6794
8.2731 | -103.6859
8.2767 | -103.699
8.2804 | -103.7185
8.2841 | -103.7426
8.2877 | -103.7637
8.2914 | -103.7847
8.2951 | -103.8036
8.2987 | -103.8122
8.3024 | -103.8229
8.3061 | -103.8362
8.3097 | -103.8431
8.3134 | -103.851
8.3171 | -103.8532
8.3207 | -103.8554
8.3244 | -103.8526
8.3281 | -103.8515
8.3317 | -103.8543
8.3354 | -103.8566
8.3391 | -103.8606
8.3427 | -103.8663
8.3464 | -103.8785
8.3501 | -103.8961
8.3537 | -103.9188
8.3574 | -103.9426
8.3611 | -103.9644
8.3647 | -103.9887
8.3684 | -104.0155
8.372 | -104.0387
8.3757 | -104.0494
8.3794 | -104.0542
8.383 | -104.0601
8.3867 | -104.0638
8.3904 | -104.0591
8.394 | -104.0449
8.3977 | -104.0232
8.4014 | -104.0002
8.405 | -103.9791
8.4087 | -103.9601
8.4124 | -103.9353
8.416 | -103.9099
8.4197 | -103.8912
8.4234 | -103.8728
8.427 | -103.8545
8.4307 | -103.8437
8.4344 | -103.8376
8.438 | -103.8402
8.4417 | -103.8463
8.4454 | -103.8578
8.449 | -103.8683
8.4527 | -103.8757
8.4564 | -103.8842
8.46 | -103.8817
8.4637 | -103.8882
8.4674 | -103.8929
8.471 | -103.8977
8.4747 | -103.8953
8.4784 | -103.8838
8.482 | -103.8728
8.4857 | -103.8649
8.4893 | -103.8528
8.493 | -103.8394
8.4967 | -103.83
8.5003 | -103.8274
8.504 | -103.8198
8.5077 | -103.8084
8.5113 | -103.7997
8.515 | -103.7977
8.5187 | -103.8033
8.5223 | -103.8112
8.526 | -103.8203
8.5297 | -103.8346
8.5333 | -103.8522
8.537 | -103.8725
8.5407 | -103.9019
8.5443 | -103.9331
8.548 | -103.9652
8.5517 | -103.9985
8.5553 | -104.0334
8.559 | -104.0675
8.5627 | -104.0961
8.5663 | -104.1224
8.57 | -104.1453
8.5737 | -104.1642
8.5773 | -104.1752
8.581 | -104.1809
8.5847 | -104.1807
8.5883 | -104.176
8.592 | -104.1635
8.5956 | -104.1429
8.5993 | -104.1177
8.603 | -104.0978
8.6066 | -104.081
8.6103 | -104.0621
8.614 | -104.0371
8.6176 | -104.0154
8.6213 | -104.0012
8.625 | -103.9887
8.6286 | -103.9729
8.6323 | -103.9655
8.636 | -103.9631
8.6396 | -103.9616
8.6433 | -103.9592
8.647 | -103.9565
8.6506 | -103.9563
8.6543 | -103.9598
8.658 | -103.9658
8.6616 | -103.9738
8.6653 | -103.9822
8.669 | -103.9847
8.6726 | -103.9881
8.6763 | -103.9884
8.68 | -103.9884
8.6836 | -103.9901
8.6873 | -103.9914
8.691 | -103.9926
8.6946 | -103.9957
8.6983 | -104.0012
8.7019 | -104.006
8.7056 | -104.0139
8.7093 | -104.0235
8.7129 | -104.0347
8.7166 | -104.0459
8.7203 | -104.0603
8.7239 | -104.08
8.7276 | -104.1
8.7313 | -104.1158
8.7349 | -104.1293
8.7386 | -104.1389
8.7423 | -104.1565
8.7459 | -104.1726
8.7496 | -104.1805
8.7533 | -104.1865
8.7569 | -104.1928
8.7606 | -104.1965
8.7643 | -104.1974
8.7679 | -104.1933
8.7716 | -104.1904
8.7753 | -104.1877
8.7789 | -104.1797
8.7826 | -104.1663
8.7863 | -104.1463
8.7899 | -104.1239
8.7936 | -104.1066
8.7973 | -104.0873
8.8009 | -104.0628
8.8046 | -104.0422
8.8082 | -104.0222
8.8119 | -103.9995
8.8156 | -103.9795
8.8192 | -103.9599
8.8229 | -103.9392
8.8266 | -103.9232
8.8302 | -103.9117
8.8339 | -103.9033
8.8376 | -103.8933
8.8412 | -103.8856
8.8449 | -103.8818
8.8486 | -103.8789
8.8522 | -103.8777
8.8559 | -103.8822
8.8596 | -103.892
8.8632 | -103.9006
8.8669 | -103.9113
8.8706 | -103.916
8.8742 | -103.9231
8.8779 | -103.9321
8.8816 | -103.939
8.8852 | -103.9443
8.8889 | -103.9452
8.8926 | -103.9466
8.8962 | -103.9529
8.8999 | -103.9535
8.9036 | -103.9462
8.9072 | -103.9479
8.9109 | -103.9579
8.9145 | -103.9711
8.9182 | -103.9851
8.9219 | -103.9976
8.9255 | -104.0142
8.9292 | -104.0336
8.9329 | -104.0505
8.9365 | -104.0666
8.9402 | -104.0837
8.9439 | -104.0997
8.9475 | -104.1152
8.9512 | -104.1277
8.9549 | -104.1348
8.9585 | -104.1392
8.9622 | -104.1433
8.9659 | -104.1483
8.9695 | -104.1521
8.9732 | -104.1521
8.9769 | -104.1537
8.9805 | -104.159
8.9842 | -104.162
8.9879 | -104.1865
8.9915 | -104.2044
8.9952 | -104.23
8.9989 | -104.2579
9.0025 | -104.2925
9.0062 | -104.3296
9.0099 | -104.3664
9.0135 | -104.4064
9.0172 | -104.4527
9.0208 | -104.4985
9.0245 | -104.5365
9.0282 | -104.5699
9.0318 | -104.602
9.0355 | -104.6333
9.0392 | -104.6501
9.0428 | -104.6569
9.0465 | -104.6596
9.0502 | -104.6545
9.0538 | -104.6393
9.0575 | -104.6188
9.0612 | -104.592
9.0648 | -104.5639
9.0685 | -104.5301
9.0722 | -104.4934
9.0758 | -104.452
9.0795 | -104.4124
9.0832 | -104.3764
9.0868 | -104.3395
9.0905 | -104.3046
9.0942 | -104.2747
9.0978 | -104.2502
9.1015 | -104.2341
9.1052 | -104.2234
9.1088 | -104.2124
9.1125 | -104.2044
9.1162 | -104.2043
9.1198 | -104.2095
9.1235 | -104.2153
9.1271 | -104.2244
9.1308 | -104.2405
9.1345 | -104.2607
9.1381 | -104.2817
9.1418 | -104.3017
9.1455 | -104.3229
9.1491 | -104.3419
9.1528 | -104.3597
9.1565 | -104.3741
9.1601 | -104.3916
9.1638 | -104.4077
9.1675 | -104.4243
9.1711 | -104.4359
9.1748 | -104.4501
9.1785 | -104.4588
9.1821 | -104.4661
9.1858 | -104.476
9.1895 | -104.4854
9.1931 | -104.4898
9.1968 | -104.4947
9.2005 | -104.4988
9.2041 | -104.5026
9.2078 | -104.5057
9.2115 | -104.5083
9.2151 | -104.5075
9.2188 | -104.5013
9.2225 | -104.4935
9.2261 | -104.4879
9.2298 | -104.481
9.2334 | -104.4765
9.2371 | -104.4705
9.2408 | -104.4683
9.2444 | -104.472
9.2481 | -104.479
9.2518 | -104.4866
9.2554 | -104.4992
9.2591 | -104.5173
9.2628 | -104.5362
9.2664 | -104.5544
9.2701 | -104.5713
9.2738 | -104.5906
9.2774 | -104.6128
9.2811 | -104.6355
9.2848 | -104.6528
9.2884 | -104.663
9.2921 | -104.6698
9.2958 | -104.6697
9.2994 | -104.6599
9.3031 | -104.6402
9.3068 | -104.6204
9.3104 | -104.5899
9.3141 | -104.5572
9.3178 | -104.5235
9.3214 | -104.4876
9.3251 | -104.4528
9.3288 | -104.423
9.3324 | -104.3974
9.3361 | -104.3772
9.3397 | -104.3616
9.3434 | -104.3554
9.3471 | -104.36
9.3507 | -104.3673
9.3544 | -104.3847
9.3581 | -104.4134
9.3617 | -104.4433
9.3654 | -104.4757
9.3691 | -104.5068
9.3727 | -104.5399
9.3764 | -104.5742
9.3801 | -104.6046
9.3837 | -104.6323
9.3874 | -104.6573
9.3911 | -104.6742
9.3947 | -104.6835
9.3984 | -104.6884
9.4021 | -104.6902
9.4057 | -104.6887
9.4094 | -104.6765
9.4131 | -104.6625
9.4167 | -104.6425
9.4204 | -104.6218
9.4241 | -104.5987
9.4277 | -104.5785
9.4314 | -104.5597
9.4351 | -104.5415
9.4387 | -104.5294
9.4424 | -104.5223
9.446 | -104.5186
9.4497 | -104.5191
9.4534 | -104.5194
9.457 | -104.5241
9.4607 | -104.534
9.4644 | -104.5456
9.468 | -104.5589
9.4717 | -104.5695
9.4754 | -104.5863
9.479 | -104.6004
9.4827 | -104.6134
9.4864 | -104.6305
9.49 | -104.6475
9.4937 | -104.6567
9.4974 | -104.6662
9.501 | -104.6765
9.5047 | -104.6876
9.5084 | -104.6957
9.512 | -104.7029
9.5157 | -104.7094
9.5194 | -104.715
9.523 | -104.7181
9.5267 | -104.7159
9.5304 | -104.7099
9.534 | -104.7019
9.5377 | -104.6913
9.5414 | -104.6761
9.545 | -104.655
9.5487 | -104.6298
9.5523 | -104.5949
9.556 | -104.5618
9.5597 | -104.5264
9.5633 | -104.483
9.567 | -104.4365
9.5707 | -104.3918
9.5743 | -104.3494
9.578 | -104.3041
9.5817 | -104.2594
9.5853 | -104.2199
9.589 | -104.1852
9.5927 | -104.1535
9.5963 | -104.1257
9.6 | -104.1037
9.6037 | -104.0874
9.6073 | -104.0755
9.611 | -104.0709
9.6147 | -104.0715
9.6183 | -104.0744
9.622 | -104.0803
9.6257 | -104.0934
9.6293 | -104.11
9.633 | -104.1336
9.6367 | -104.1579
9.6403 | -104.1827
9.644 | -104.2108
9.6477 | -104.2392
9.6513 | -104.2652
9.655 | -104.292
9.6586 | -104.3158
9.6623 | -104.3404
9.666 | -104.3604
9.6696 | -104.3807
9.6733 | -104.4013
9.677 | -104.4164
9.6806 | -104.4261
9.6843 | -104.4327
9.688 | -104.4365
9.6916 | -104.4369
9.6953 | -104.4346
9.699 | -104.4308
9.7026 | -104.4297
9.7063 | -104.4316
9.71 | -104.4313
9.7136 | -104.4313
9.7173 | -104.4338
9.721 | -104.4408
9.7246 | -104.4502
9.7283 | -104.4607
9.732 | -104.4717
9.7356 | -104.49
9.7393 | -104.5107
9.743 | -104.5317
9.7466 | -104.5557
9.7503 | -104.5829
9.754 | -104.6093
9.7576 | -104.6318
9.7613 | -104.6613
9.7649 | -104.6927
9.7686 | -104.7204
9.7723 | -104.7502
9.7759 | -104.7773
9.7796 | -104.8025
9.7833 | -104.8278
9.7869 | -104.8489
9.7906 | -104.8733
9.7943 | -104.8986
9.7979 | -104.919
9.8016 | -104.9419
9.8053 | -104.9567
9.8089 | -104.9721
9.8126 | -104.9839
9.8163 | -104.9877
9.8199 | -104.9934
9.8236 | -104.9936
9.8273 | -104.9921
9.8309 | -104.9909
9.8346 | -104.9881
9.8383 | -104.9816
9.8419 | -104.9663
9.8456 | -104.9499
9.8493 | -104.9374
9.8529 | -104.929
9.8566 | -104.9198
9.8603 | -104.9101
9.8639 | -104.897
9.8676 | -104.8898
9.8712 | -104.8852
9.8749 | -104.883
9.8786 | -104.8769
9.8822 | -104.8696
9.8859 | -104.8617
9.8896 | -104.8557
9.8932 | -104.8485
9.8969 | -104.8417
9.9006 | -104.8351
9.9042 | -104.8222
9.9079 | -104.8046
9.9116 | -104.7866
9.9152 | -104.761
9.9189 | -104.7325
9.9226 | -104.7033
9.9262 | -104.6645
9.9299 | -104.6247
9.9336 | -104.5774
9.9372 | -104.5261
9.9409 | -104.4742
9.9446 | -104.4206
9.9482 | -104.3729
9.9519 | -104.3204
9.9556 | -104.2709
9.9592 | -104.234
9.9629 | -104.1995
9.9666 | -104.177
9.9702 | -104.1604
9.9739 | -104.1481
9.9775 | -104.1469
9.9812 | -104.1548
9.9849 | -104.1715
9.9885 | -104.1926
9.9922 | -104.2203
9.9959 | -104.2482
9.9995 | -104.2811
10.0032 | -104.3169
10.0069 | -104.3579
10.0105 | -104.3913
10.0142 | -104.4279
10.0179 | -104.4641
10.0215 | -104.4953
10.0252 | -104.5195
10.0289 | -104.5413
10.0325 | -104.5591
10.0362 | -104.5728
10.0399 | -104.5771
10.0435 | -104.5763
10.0472 | -104.5659
10.0509 | -104.5639
10.0545 | -104.5592
10.0582 | -104.5519
10.0619 | -104.545
10.0655 | -104.5386
10.0692 | -104.5263
10.0729 | -104.5136
10.0765 | -104.5001
10.0802 | -104.4887
10.0838 | -104.4783
10.0875 | -104.469
10.0912 | -104.4631
10.0948 | -104.4599
10.0985 | -104.4639
10.1022 | -104.4667
10.1058 | -104.468
10.1095 | -104.4731
10.1132 | -104.4797
10.1168 | -104.4848
10.1205 | -104.4951
10.1242 | -104.5066
10.1278 | -104.5151
10.1315 | -104.5269
10.1352 | -104.5372
10.1388 | -104.5434
10.1425 | -104.5546
10.1462 | -104.5621
10.1498 | -104.5704
10.1535 | -104.5784
10.1572 | -104.5924
10.1608 | -104.604
10.1645 | -104.6139
10.1682 | -104.6323
10.1718 | -104.6534
10.1755 | -104.6759
10.1792 | -104.697
10.1828 | -104.7193
10.1865 | -104.7427
10.1901 | -104.7673
10.1938 | -104.7867
10.1975 | -104.8039
10.2011 | -104.8238
10.2048 | -104.8373
10.2085 | -104.846
10.2121 | -104.8529
10.2158 | -104.8555
10.2195 | -104.8569
10.2231 | -104.8552
10.2268 | -104.8447
10.2305 | -104.8311
10.2341 | -104.8143
10.2378 | -104.7954
10.2415 | -104.7737
10.2451 | -104.7494
10.2488 | -104.7233
10.2525 | -104.6997
10.2561 | -104.6742
10.2598 | -104.6442
10.2635 | -104.6182
10.2671 | -104.6025
10.2708 | -104.5862
10.2745 | -104.5694
10.2781 | -104.5593
10.2818 | -104.5484
10.2855 | -104.5405
10.2891 | -104.5364
10.2928 | -104.532
10.2964 | -104.5304
10.3001 | -104.5318
10.3038 | -104.5355
10.3074 | -104.5367
10.3111 | -104.5398
10.3148 | -104.5441
10.3184 | -104.548
10.3221 | -104.5569
10.3258 | -104.5648
10.3294 | -104.5738
10.3331 | -104.5834
10.3368 | -104.5968
10.3404 | -104.6097
10.3441 | -104.6223
10.3478 | -104.6345
10.3514 | -104.6478
10.3551 | -104.6677
10.3588 | -104.6879
10.3624 | -104.7043
10.3661 | -104.7184
10.3698 | -104.7348
10.3734 | -104.7523
10.3771 | -104.7703
10.3808 | -104.7848
10.3844 | -104.7989
10.3881 | -104.8121
10.3918 | -104.8272
10.3954 | -104.8398
10.3991 | -104.8488
10.4027 | -104.8575
10.4064 | -104.8679
10.4101 | -104.8822
10.4137 | -104.8917
10.4174 | -104.8948
10.4211 | -104.8955
10.4247 | -104.8984
10.4284 | -104.9041
10.4321 | -104.9078
10.4357 | -104.9087
10.4394 | -104.9133
10.4431 | -104.9169
10.4467 | -104.917
10.4504 | -104.914
10.4541 | -104.909
10.4577 | -104.9054
10.4614 | -104.8997
10.4651 | -104.8924
10.4687 | -104.8816
10.4724 | -104.8667
10.4761 | -104.8501
10.4797 | -104.8321
10.4834 | -104.8151
10.4871 | -104.7926
10.4907 | -104.7662
10.4944 | -104.7431
10.4981 | -104.7209
10.5017 | -104.699
10.5054 | -104.6785
10.509 | -104.6602
10.5127 | -104.6485
10.5164 | -104.6411
10.52 | -104.6328
10.5237 | -104.6289
10.5274 | -104.6305
10.531 | -104.6369
10.5347 | -104.6489
10.5384 | -104.6675
10.542 | -104.6872
10.5457 | -104.713
10.5494 | -104.7494
10.553 | -104.7903
10.5567 | -104.8381
10.5604 | -104.8903
10.564 | -104.9421
10.5677 | -104.9979
10.5714 | -105.0586
10.575 | -105.1114
10.5787 | -105.161
10.5824 | -105.2052
10.586 | -105.2499
10.5897 | -105.2905
10.5934 | -105.3183
10.597 | -105.3355
10.6007 | -105.3488
10.6044 | -105.3535
10.608 | -105.3496
10.6117 | -105.3385
10.6153 | -105.3197
10.619 | -105.2936
10.6227 | -105.2624
10.6263 | -105.2282
10.63 | -105.1956
10.6337 | -105.1578
10.6373 | -105.1192
10.641 | -105.0849
10.6447 | -105.0495
10.6483 | -105.0166
10.652 | -104.9851
10.6557 | -104.9546
10.6593 | -104.9243
10.663 | -104.8964
10.6667 | -104.8753
10.6703 | -104.858
10.674 | -104.8411
10.6777 | -104.8214
10.6813 | -104.8032
10.685 | -104.7934
10.6887 | -104.7852
10.6923 | -104.7758
10.696 | -104.7683
10.6997 | -104.7622
10.7033 | -104.758
10.707 | -104.7523
10.7107 | -104.7426
10.7143 | -104.7293
10.718 | -104.7213
10.7216 | -104.716
10.7253 | -104.7085
10.729 | -104.699
10.7326 | -104.6902
10.7363 | -104.685
10.74 | -104.6797
10.7436 | -104.6751
10.7473 | -104.671
10.751 | -104.6709
10.7546 | -104.6727
10.7583 | -104.6734
10.762 | -104.6751
10.7656 | -104.6797
10.7693 | -104.686
10.773 | -104.694
10.7766 | -104.7048
10.7803 | -104.7167
10.784 | -104.7289
10.7876 | -104.7411
10.7913 | -104.7591
10.795 | -104.7771
10.7986 | -104.7959
10.8023 | -104.8152
10.806 | -104.8307
10.8096 | -104.847
10.8133 | -104.864
10.817 | -104.883
10.8206 | -104.8992
10.8243 | -104.9149
10.8279 | -104.9324
10.8316 | -104.9539
10.8353 | -104.9757
10.8389 | -104.997
10.8426 | -105.0105
10.8463 | -105.0196
10.8499 | -105.0282
10.8536 | -105.0343
10.8573 | -105.042
10.8609 | -105.0466
10.8646 | -105.0486
10.8683 | -105.0534
10.8719 | -105.0583
10.8756 | -105.0579
10.8793 | -105.0514
10.8829 | -105.0421
10.8866 | -105.0308
10.8903 | -105.0196
10.8939 | -105.0068
10.8976 | -104.9907
10.9013 | -104.9687
10.9049 | -104.951
10.9086 | -104.9358
10.9123 | -104.9215
10.9159 | -104.9086
10.9196 | -104.8999
10.9233 | -104.8902
10.9269 | -104.8851
10.9306 | -104.8879
10.9343 | -104.8909
10.9379 | -104.8972
10.9416 | -104.9101
10.9452 | -104.9252
10.9489 | -104.9407
10.9526 | -104.9628
10.9562 | -104.9884
10.9599 | -105.0164
10.9636 | -105.0439
10.9672 | -105.0715
10.9709 | -105.1002
10.9746 | -105.1274
10.9782 | -105.1506
10.9819 | -105.1701
10.9856 | -105.1819
10.9892 | -105.1925
10.9929 | -105.2039
10.9966 | -105.2073
11.0002 | -105.209
11.0039 | -105.211
11.0076 | -105.209
11.0112 | -105.202
11.0149 | -105.1896
11.0186 | -105.1706
11.0222 | -105.1523
11.0259 | -105.1278
11.0296 | -105.0979
11.0332 | -105.0665
11.0369 | -105.0309
11.0405 | -104.996
11.0442 | -104.961
11.0479 | -104.924
11.0515 | -104.8873
11.0552 | -104.8556
11.0589 | -104.8282
11.0625 | -104.7973
11.0662 | -104.7667
11.0699 | -104.7419
11.0735 | -104.72
11.0772 | -104.6989
11.0809 | -104.6828
11.0845 | -104.6702
11.0882 | -104.6585
11.0919 | -104.6512
11.0955 | -104.6453
11.0992 | -104.641
11.1029 | -104.6397
11.1065 | -104.6366
11.1102 | -104.636
11.1139 | -104.6321
11.1175 | -104.626
11.1212 | -104.6218
11.1249 | -104.6263
11.1285 | -104.6299
11.1322 | -104.6303
11.1359 | -104.633
11.1395 | -104.6394
11.1432 | -104.645
11.1469 | -104.6552
11.1505 | -104.6652
11.1542 | -104.6751
11.1578 | -104.688
11.1615 | -104.7018
11.1652 | -104.7178
11.1688 | -104.7353
11.1725 | -104.7522
11.1762 | -104.7687
11.1798 | -104.7894
11.1835 | -104.811
11.1872 | -104.8345
11.1908 | -104.8621
11.1945 | -104.8863
11.1982 | -104.907
11.2018 | -104.9277
11.2055 | -104.9437
11.2092 | -104.9558
11.2128 | -104.9655
11.2165 | -104.9688
11.2202 | -104.9695
11.2238 | -104.9701
11.2275 | -104.9697
11.2312 | -104.9653
11.2348 | -104.9567
11.2385 | -104.9465
11.2422 | -104.9366
11.2458 | -104.9266
11.2495 | -104.9165
11.2531 | -104.9063
11.2568 | -104.898
11.2605 | -104.8937
11.2641 | -104.8947
11.2678 | -104.8984
11.2715 | -104.9025
11.2751 | -104.9093
11.2788 | -104.9225
11.2825 | -104.9408
11.2861 | -104.9582
11.2898 | -104.9711
11.2935 | -104.9897
11.2971 | -105.01
11.3008 | -105.029
11.3045 | -105.0482
11.3081 | -105.0684
11.3118 | -105.0885
11.3155 | -105.1056
11.3191 | -105.1192
11.3228 | -105.1342
11.3265 | -105.147
11.3301 | -105.1546
11.3338 | -105.1614
11.3375 | -105.1659
11.3411 | -105.1662
11.3448 | -105.1604
11.3485 | -105.1532
11.3521 | -105.1413
11.3558 | -105.1263
11.3595 | -105.113
11.3631 | -105.1063
11.3668 | -105.0945
11.3704 | -105.0792
11.3741 | -105.0651
11.3778 | -105.0519
11.3814 | -105.0365
11.3851 | -105.0202
11.3888 | -105.0031
11.3924 | -104.9882
11.3961 | -104.9752
11.3998 | -104.9649
11.4034 | -104.9516
11.4071 | -104.9365
11.4108 | -104.924
11.4144 | -104.9134
11.4181 | -104.9036
11.4218 | -104.8945
11.4254 | -104.882
11.4291 | -104.867
11.4328 | -104.8541
11.4364 | -104.8445
11.4401 | -104.8354
11.4438 | -104.8263
11.4474 | -104.8178
11.4511 | -104.8139
11.4548 | -104.8117
11.4584 | -104.8073
11.4621 | -104.8057
11.4657 | -104.8061
11.4694 | -104.808
11.4731 | -104.8095
11.4767 | -104.8123
11.4804 | -104.8128
11.4841 | -104.8137
11.4877 | -104.8149
11.4914 | -104.8181
11.4951 | -104.8205
11.4987 | -104.8238
11.5024 | -104.8266
11.5061 | -104.8282
11.5097 | -104.83
11.5134 | -104.8324
11.5171 | -104.8355
11.5207 | -104.8373
11.5244 | -104.8355
11.5281 | -104.8336
11.5317 | -104.8331
11.5354 | -104.8348
11.5391 | -104.8351
11.5427 | -104.8357
11.5464 | -104.8358
11.5501 | -104.8355
11.5537 | -104.8339
11.5574 | -104.8281
11.5611 | -104.822
11.5647 | -104.8184
11.5684 | -104.8185
11.5721 | -104.8203
11.5757 | -104.8186
11.5794 | -104.8196
11.583 | -104.8251
11.5867 | -104.8296
11.5904 | -104.8373
11.594 | -104.8517
11.5977 | -104.8691
11.6014 | -104.8876
11.605 | -104.9075
11.6087 | -104.9283
11.6124 | -104.9531
11.616 | -104.9785
11.6197 | -105.0054
11.6234 | -105.0348
11.627 | -105.0644
11.6307 | -105.0919
11.6344 | -105.1199
11.638 | -105.1461
11.6417 | -105.1677
11.6454 | -105.189
11.649 | -105.2124
11.6527 | -105.231
11.6564 | -105.2466
11.66 | -105.2593
11.6637 | -105.2735
11.6674 | -105.282
11.671 | -105.2893
11.6747 | -105.2957
11.6784 | -105.297
11.682 | -105.2975
11.6857 | -105.2999
11.6893 | -105.301
11.693 | -105.298
11.6967 | -105.2928
11.7003 | -105.2881
11.704 | -105.283
11.7077 | -105.2775
11.7113 | -105.2718
11.715 | -105.2666
11.7187 | -105.2637
11.7223 | -105.2625
11.726 | -105.2657
11.7297 | -105.2674
11.7333 | -105.2686
11.737 | -105.2699
11.7407 | -105.2714
11.7443 | -105.2736
11.748 | -105.2744
11.7517 | -105.2722
11.7553 | -105.2717
11.759 | -105.2711
11.7627 | -105.2683
11.7663 | -105.2609
11.77 | -105.2555
11.7737 | -105.2508
11.7773 | -105.2439
11.781 | -105.232
11.7847 | -105.2193
11.7883 | -105.2071
11.792 | -105.1965
11.7956 | -105.1865
11.7993 | -105.1767
11.803 | -105.165
11.8066 | -105.1544
11.8103 | -105.1481
11.814 | -105.1465
11.8176 | -105.145
11.8213 | -105.1465
11.825 | -105.1524
11.8286 | -105.1605
11.8323 | -105.166
11.836 | -105.173
11.8396 | -105.1839
11.8433 | -105.1988
11.847 | -105.2155
11.8506 | -105.2353
11.8543 | -105.2492
11.858 | -105.2607
11.8616 | -105.2681
11.8653 | -105.2782
11.869 | -105.2819
11.8726 | -105.2843
11.8763 | -105.2868
11.88 | -105.2869
11.8836 | -105.281
11.8873 | -105.2717
11.891 | -105.2568
11.8946 | -105.2398
11.8983 | -105.2208
11.9019 | -105.2024
11.9056 | -105.1852
11.9093 | -105.1706
11.9129 | -105.1562
11.9166 | -105.1451
11.9203 | -105.1361
11.9239 | -105.1288
11.9276 | -105.1223
11.9313 | -105.1177
11.9349 | -105.114
11.9386 | -105.1119
11.9423 | -105.1086
11.9459 | -105.1106
11.9496 | -105.1162
11.9533 | -105.1225
11.9569 | -105.133
11.9606 | -105.1432
11.9643 | -105.1524
11.9679 | -105.1608
11.9716 | -105.1692
11.9753 | -105.1835
11.9789 | -105.1935
11.9826 | -105.1993
11.9863 | -105.2071
11.9899 | -105.2145
11.9936 | -105.2189
11.9973 | -105.2241
12.0009 | -105.2259
12.0046 | -105.2317
12.0082 | -105.2345
12.0119 | -105.2357
12.0156 | -105.2351
12.0192 | -105.233
12.0229 | -105.2293
12.0266 | -105.2294
12.0302 | -105.2286
12.0339 | -105.2284
12.0376 | -105.2274
12.0412 | -105.2267
12.0449 | -105.2236
12.0486 | -105.2183
12.0522 | -105.2139
12.0559 | -105.2144
12.0596 | -105.214
12.0632 | -105.213
12.0669 | -105.2113
12.0706 | -105.2065
12.0742 | -105.2049
12.0779 | -105.202
12.0816 | -105.1996
12.0852 | -105.1981
12.0889 | -105.1977
12.0926 | -105.2005
12.0962 | -105.2048
12.0999 | -105.2098
12.1036 | -105.2151
12.1072 | -105.2203
12.1109 | -105.2249
12.1145 | -105.2289
12.1182 | -105.2305
12.1219 | -105.2347
12.1255 | -105.24
12.1292 | -105.2432
12.1329 | -105.2465
12.1365 | -105.2494
12.1402 | -105.2521
12.1439 | -105.2534
12.1475 | -105.2477
12.1512 | -105.2448
12.1549 | -105.2412
12.1585 | -105.2366
12.1622 | -105.2309
12.1659 | -105.2263
12.1695 | -105.2221
12.1732 | -105.216
12.1769 | -105.2105
12.1805 | -105.2037
12.1842 | -105.2
12.1879 | -105.1995
12.1915 | -105.1988
12.1952 | -105.1969
12.1989 | -105.1962
12.2025 | -105.1958
12.2062 | -105.1989
12.2099 | -105.2057
12.2135 | -105.2057
12.2172 | -105.2073
12.2208 | -105.2082
12.2245 | -105.2091
12.2282 | -105.213
12.2318 | -105.2124
12.2355 | -105.2143
12.2392 | -105.2147
12.2428 | -105.2141
12.2465 | -105.2144
12.2502 | -105.212
12.2538 | -105.2057
12.2575 | -105.2011
12.2612 | -105.1961
12.2648 | -105.1877
12.2685 | -105.1752
12.2722 | -105.1647
12.2758 | -105.1568
12.2795 | -105.1472
12.2832 | -105.1377
12.2868 | -105.1295
12.2905 | -105.1206
12.2942 | -105.1094
12.2978 | -105.093
12.3015 | -105.082
12.3052 | -105.0739
12.3088 | -105.0691
12.3125 | -105.0665
12.3162 | -105.0656
12.3198 | -105.0686
12.3235 | -105.0762
12.3271 | -105.0795
12.3308 | -105.0853
12.3345 | -105.0879
12.3381 | -105.0906
12.3418 | -105.0935
12.3455 | -105.097
12.3491 | -105.0992
12.3528 | -105.0983
12.3565 | -105.0971
12.3601 | -105.0939
12.3638 | -105.089
12.3675 | -105.0845
12.3711 | -105.0804
12.3748 | -105.0752
12.3785 | -105.0693
12.3821 | -105.0632
12.3858 | -105.0569
12.3895 | -105.0486
12.3931 | -105.0466
12.3968 | -105.0389
12.4005 | -105.033
12.4041 | -105.029
12.4078 | -105.0211
12.4115 | -105.0137
12.4151 | -105.0092
12.4188 | -105.007
12.4225 | -105.0065
12.4261 | -105.0063
12.4298 | -105.0074
12.4334 | -105.0088
12.4371 | -105.0106
12.4408 | -105.0137
12.4444 | -105.0164
12.4481 | -105.0205
12.4518 | -105.0238
12.4554 | -105.0247
12.4591 | -105.0256
12.4628 | -105.0265
12.4664 | -105.0277
12.4701 | -105.0314
12.4738 | -105.036
12.4774 | -105.0451
12.4811 | -105.0496
12.4848 | -105.0497
12.4884 | -105.0461
12.4921 | -105.045
12.4958 | -105.0473
12.4994 | -105.0468
12.5031 | -105.0477
12.5068 | -105.0507
12.5104 | -105.0531
12.5141 | -105.0577
12.5178 | -105.0595
12.5214 | -105.0617
12.5251 | -105.0669
12.5288 | -105.0755
12.5324 | -105.0861
12.5361 | -105.0953
12.5397 | -105.1037
12.5434 | -105.1095
12.5471 | -105.1146
12.5507 | -105.1185
12.5544 | -105.1249
12.5581 | -105.1323
12.5617 | -105.1386
12.5654 | -105.1455
12.5691 | -105.1517
12.5727 | -105.1562
12.5764 | -105.1652
12.5801 | -105.1702
12.5837 | -105.1728
12.5874 | -105.1776
12.5911 | -105.1825
12.5947 | -105.1828
12.5984 | -105.1827
12.6021 | -105.183
12.6057 | -105.1821
12.6094 | -105.183
12.6131 | -105.1854
12.6167 | -105.1823
12.6204 | -105.1846
12.6241 | -105.191
12.6277 | -105.1943
12.6314 | -105.1975
12.6351 | -105.2013
12.6387 | -105.2059
12.6424 | -105.2116
12.646 | -105.2203
12.6497 | -105.2252
12.6534 | -105.2304
12.657 | -105.2365
12.6607 | -105.2439
12.6644 | -105.2523
12.668 | -105.2642
12.6717 | -105.2749
12.6754 | -105.2813
12.679 | -105.2906
12.6827 | -105.3015
12.6864 | -105.3111
12.69 | -105.3141
12.6937 | -105.3202
12.6974 | -105.3252
12.701 | -105.3315
12.7047 | -105.3351
12.7084 | -105.3371
12.712 | -105.3388
12.7157 | -105.3428
12.7194 | -105.3465
12.723 | -105.3499
12.7267 | -105.3538
12.7304 | -105.3542
12.734 | -105.3545
12.7377 | -105.3545
12.7414 | -105.3559
12.745 | -105.359
12.7487 | -105.3564
12.7523 | -105.3551
12.756 | -105.352
12.7597 | -105.3448
12.7633 | -105.339
12.767 | -105.3335
12.7707 | -105.3258
12.7743 | -105.3211
12.778 | -105.3128
12.7817 | -105.3027
12.7853 | -105.2923
12.789 | -105.2796
12.7927 | -105.2663
12.7963 | -105.2519
12.8 | -105.2405
12.8037 | -105.2259
12.8073 | -105.2122
12.811 | -105.2025
12.8147 | -105.1921
12.8183 | -105.183
12.822 | -105.1765
12.8257 | -105.1729
12.8293 | -105.1675
12.833 | -105.1646
12.8367 | -105.1634
12.8403 | -105.1629
12.844 | -105.165
12.8477 | -105.1697
12.8513 | -105.1761
12.855 | -105.1869
12.8586 | -105.1998
12.8623 | -105.2135
12.866 | -105.227
12.8696 | -105.2434
12.8733 | -105.2635
12.877 | -105.2813
12.8806 | -105.3007
12.8843 | -105.3186
12.888 | -105.338
12.8916 | -105.3566
12.8953 | -105.3717
12.899 | -105.3854
12.9026 | -105.3945
12.9063 | -105.4008
12.91 | -105.409
12.9136 | -105.4137
12.9173 | -105.4197
12.921 | -105.4234
12.9246 | -105.4212
12.9283 | -105.4142
12.932 | -105.4062
12.9356 | -105.4019
12.9393 | -105.3943
12.943 | -105.3863
12.9466 | -105.3779
12.9503 | -105.3656
12.954 | -105.3507
12.9576 | -105.3345
12.9613 | -105.3178
12.9649 | -105.3064
12.9686 | -105.2926
12.9723 | -105.2776
12.9759 | -105.2663
12.9796 | -105.251
12.9833 | -105.2398
12.9869 | -105.2295
12.9906 | -105.2177
12.9943 | -105.204
12.9979 | -105.1931
13.0016 | -105.1851
13.0053 | -105.1772
13.0089 | -105.1691
13.0126 | -105.1587
13.0163 | -105.1512
13.0199 | -105.1479
13.0236 | -105.1439
13.0273 | -105.1401
13.0309 | -105.1393
13.0346 | -105.1399
13.0383 | -105.1391
13.0419 | -105.1347
13.0456 | -105.1345
13.0493 | -105.1345
13.0529 | -105.1359
13.0566 | -105.1391
13.0603 | -105.1409
13.0639 | -105.1469
13.0676 | -105.1551
13.0712 | -105.1644
13.0749 | -105.1741
13.0786 | -105.1845
13.0822 | -105.1931
13.0859 | -105.2031
13.0896 | -105.2191
13.0932 | -105.2382
13.0969 | -105.2538
13.1006 | -105.2685
13.1042 | -105.28
13.1079 | -105.29
13.1116 | -105.3015
13.1152 | -105.3125
13.1189 | -105.3221
13.1226 | -105.3285
13.1262 | -105.3379
13.1299 | -105.3457
13.1336 | -105.3517
13.1372 | -105.3586
13.1409 | -105.3656
13.1446 | -105.3665
13.1482 | -105.3617
13.1519 | -105.3573
13.1556 | -105.357
13.1592 | -105.3551
13.1629 | -105.352
13.1666 | -105.35
13.1702 | -105.3402
13.1739 | -105.3279
13.1775 | -105.3174
13.1812 | -105.3096
13.1849 | -105.2989
13.1885 | -105.2862
13.1922 | -105.2727
13.1959 | -105.2638
13.1995 | -105.2558
13.2032 | -105.2464
13.2069 | -105.2368
13.2105 | -105.2292
13.2142 | -105.2212
13.2179 | -105.2235
13.2215 | -105.2187
13.2252 | -105.2142
13.2289 | -105.2124
13.2325 | -105.2136
13.2362 | -105.2144
13.2399 | -105.2168
13.2435 | -105.2186
13.2472 | -105.2218
13.2509 | -105.2266
13.2545 | -105.2286
13.2582 | -105.2314
13.2619 | -105.2367
13.2655 | -105.2434
13.2692 | -105.251
13.2729 | -105.2631
13.2765 | -105.2719
13.2802 | -105.2753
13.2838 | -105.2781
13.2875 | -105.2808
13.2912 | -105.2859
13.2948 | -105.2918
13.2985 | -105.2977
13.3022 | -105.2997
13.3058 | -105.3003
13.3095 | -105.2986
13.3132 | -105.2975
13.3168 | -105.294
13.3205 | -105.291
13.3242 | -105.2889
13.3278 | -105.2858
13.3315 | -105.2832
13.3352 | -105.2855
13.3388 | -105.2875
13.3425 | -105.2867
13.3462 | -105.2849
13.3498 | -105.2851
13.3535 | -105.2869
13.3572 | -105.2902
13.3608 | -105.2948
13.3645 | -105.3021
13.3682 | -105.3103
13.3718 | -105.3181
13.3755 | -105.3261
13.3792 | -105.3362
13.3828 | -105.3509
13.3865 | -105.363
13.3901 | -105.3763
13.3938 | -105.3914
13.3975 | -105.4062
13.4011 | -105.4193
13.4048 | -105.4335
13.4085 | -105.4455
13.4121 | -105.4593
13.4158 | -105.4738
13.4195 | -105.4901
13.4231 | -105.503
13.4268 | -105.5118
13.4305 | -105.5169
13.4341 | -105.5237
13.4378 | -105.5308
13.4415 | -105.5364
13.4451 | -105.5391
13.4488 | -105.5408
13.4525 | -105.5409
13.4561 | -105.5407
13.4598 | -105.5343
13.4635 | -105.5278
13.4671 | -105.518
13.4708 | -105.5082
13.4745 | -105.5025
13.4781 | -105.4968
13.4818 | -105.4928
13.4855 | -105.4867
13.4891 | -105.4761
13.4928 | -105.4671
13.4964 | -105.4637
13.5001 | -105.4581
13.5038 | -105.4537
13.5074 | -105.4477
13.5111 | -105.4422
13.5148 | -105.4398
13.5184 | -105.4371
13.5221 | -105.4382
13.5258 | -105.4391
13.5294 | -105.4386
13.5331 | -105.4408
13.5368 | -105.4409
13.5404 | -105.4428
13.5441 | -105.4478
13.5478 | -105.4515
13.5514 | -105.4553
13.5551 | -105.4624
13.5588 | -105.4666
13.5624 | -105.4711
13.5661 | -105.4741
13.5698 | -105.4777
13.5734 | -105.4791
13.5771 | -105.4765
13.5808 | -105.4725
13.5844 | -105.4708
13.5881 | -105.4682
13.5918 | -105.4629
13.5954 | -105.4585
13.5991 | -105.4521
13.6027 | -105.4445
13.6064 | -105.4342
13.6101 | -105.4226
13.6137 | -105.4134
13.6174 | -105.4061
13.6211 | -105.3976
13.6247 | -105.3907
13.6284 | -105.383
13.6321 | -105.3738
13.6357 | -105.3685
13.6394 | -105.3662
13.6431 | -105.3629
13.6467 | -105.3611
13.6504 | -105.3594
13.6541 | -105.3615
13.6577 | -105.3616
13.6614 | -105.3624
13.6651 | -105.3648
13.6687 | -105.3664
13.6724 | -105.3662
13.6761 | -105.3639
13.6797 | -105.3617
13.6834 | -105.3643
13.6871 | -105.3653
13.6907 | -105.3627
13.6944 | -105.3607
13.6981 | -105.3595
13.7017 | -105.3598
13.7054 | -105.356
13.709 | -105.3481
13.7127 | -105.3424
13.7164 | -105.3346
13.72 | -105.3278
13.7237 | -105.3223
13.7274 | -105.3166
13.731 | -105.3116
13.7347 | -105.3064
13.7384 | -105.2964
13.742 | -105.2867
13.7457 | -105.2778
13.7494 | -105.2709
13.753 | -105.2616
13.7567 | -105.2509
13.7604 | -105.2413
13.764 | -105.2329
13.7677 | -105.2255
13.7714 | -105.2206
13.775 | -105.2157
13.7787 | -105.21
13.7824 | -105.2045
13.786 | -105.1992
13.7897 | -105.1946
13.7934 | -105.189
13.797 | -105.1859
13.8007 | -105.1814
13.8044 | -105.1781
13.808 | -105.1756
13.8117 | -105.1663
13.8153 | -105.1544
13.819 | -105.1473
13.8227 | -105.1399
13.8263 | -105.1313
13.83 | -105.1219
13.8337 | -105.1089
13.8373 | -105.0954
13.841 | -105.0818
13.8447 | -105.0668
13.8483 | -105.0464
13.852 | -105.0278
13.8557 | -105.0078
13.8593 | -104.9864
13.863 | -104.9651
13.8667 | -104.9402
13.8703 | -104.9153
13.874 | -104.8938
13.8777 | -104.8732
13.8813 | -104.8491
13.885 | -104.8281
13.8887 | -104.8092
13.8923 | -104.7915
13.896 | -104.7739
13.8997 | -104.7596
13.9033 | -104.7503
13.907 | -104.7428
13.9107 | -104.7365
13.9143 | -104.7321
13.918 | -104.7322
13.9216 | -104.7362
13.9253 | -104.7431
13.929 | -104.7521
13.9326 | -104.7642
13.9363 | -104.7823
13.94 | -104.8042
13.9436 | -104.8269
13.9473 | -104.8506
13.951 | -104.8767
13.9546 | -104.9049
13.9583 | -104.9357
13.962 | -104.9594
13.9656 | -104.9812
13.9693 | -105.0055
13.973 | -105.0289
13.9766 | -105.054
13.9803 | -105.0753
13.984 | -105.0908
13.9876 | -105.1078
13.9913 | -105.1223
13.995 | -105.1341
13.9986 | -105.141
14.0023 | -105.1453
14.006 | -105.1472
14.0096 | -105.1456
14.0133 | -105.145
14.017 | -105.1424
14.0206 | -105.1332
14.0243 | -105.1222
14.028 | -105.109
14.0316 | -105.0934
14.0353 | -105.0791
14.0389 | -105.0637
14.0426 | -105.0507
14.0463 | -105.0394
14.0499 | -105.0313
14.0536 | -105.0209
14.0573 | -105.0131
14.0609 | -105.0049
14.0646 | -104.9981
14.0683 | -104.9976
14.0719 | -104.9986
14.0756 | -104.9981
14.0793 | -105.0007
14.0829 | -105.0081
14.0866 | -105.0158
14.0903 | -105.0261
14.0939 | -105.0421
14.0976 | -105.0606
14.1013 | -105.0801
14.1049 | -105.1006
14.1086 | -105.1229
14.1123 | -105.145
14.1159 | -105.1669
14.1196 | -105.185
14.1233 | -105.2036
14.1269 | -105.222
14.1306 | -105.2402
14.1342 | -105.2539
14.1379 | -105.2643
14.1416 | -105.2737
14.1452 | -105.2817
14.1489 | -105.2868
14.1526 | -105.2916
14.1562 | -105.297
14.1599 | -105.2954
14.1636 | -105.2974
14.1672 | -105.3011
14.1709 | -105.3004
14.1746 | -105.2971
14.1782 | -105.2915
14.1819 | -105.2896
14.1856 | -105.2875
14.1892 | -105.2816
14.1929 | -105.2741
14.1966 | -105.2638
14.2002 | -105.2532
14.2039 | -105.2447
14.2076 | -105.2363
14.2112 | -105.2281
14.2149 | -105.2201
14.2186 | -105.2134
14.2222 | -105.2091
14.2259 | -105.2035
14.2296 | -105.1961
14.2332 | -105.1911
14.2369 | -105.1874
14.2406 | -105.1859
14.2442 | -105.1832
14.2479 | -105.1788
14.2515 | -105.1755
14.2552 | -105.1743
14.2589 | -105.175
14.2625 | -105.1789
14.2662 | -105.181
14.2699 | -105.1792
14.2735 | -105.1819
14.2772 | -105.1881
14.2809 | -105.193
14.2845 | -105.1976
14.2882 | -105.2003
14.2919 | -105.2018
14.2955 | -105.2035
14.2992 | -105.2035
14.3029 | -105.2039
14.3065 | -105.204
14.3102 | -105.203
14.3139 | -105.2034
14.3175 | -105.2048
14.3212 | -105.206
14.3249 | -105.2085
14.3285 | -105.2094
14.3322 | -105.2089
14.3359 | -105.2098
14.3395 | -105.2135
14.3432 | -105.217
14.3468 | -105.2182
14.3505 | -105.2192
14.3542 | -105.2187
14.3578 | -105.2178
14.3615 | -105.2164
14.3652 | -105.2188
14.3688 | -105.2217
14.3725 | -105.2225
14.3762 | -105.2218
14.3798 | -105.2236
14.3835 | -105.2267
14.3872 | -105.2302
14.3908 | -105.2354
14.3945 | -105.2419
14.3982 | -105.252
14.4018 | -105.2655
14.4055 | -105.2783
14.4092 | -105.2896
14.4128 | -105.3022
14.4165 | -105.3156
14.4202 | -105.3306
14.4238 | -105.3467
14.4275 | -105.3623
14.4312 | -105.3777
14.4348 | -105.3964
14.4385 | -105.417
14.4422 | -105.4315
14.4458 | -105.4474
14.4495 | -105.464
14.4532 | -105.4806
14.4568 | -105.4994
14.4605 | -105.5183
14.4641 | -105.5343
14.4678 | -105.5482
14.4715 | -105.5595
14.4751 | -105.5703
14.4788 | -105.5809
14.4825 | -105.5907
14.4861 | -105.5974
14.4898 | -105.6043
14.4935 | -105.6099
14.4971 | -105.6158
14.5008 | -105.6223
14.5045 | -105.6214
14.5081 | -105.6198
14.5118 | -105.6161
14.5155 | -105.611
14.5191 | -105.6086
14.5228 | -105.6032
14.5265 | -105.5961
14.5301 | -105.5893
14.5338 | -105.5812
14.5375 | -105.5726
14.5411 | -105.5648
14.5448 | -105.5533
14.5485 | -105.5446
14.5521 | -105.5369
14.5558 | -105.5296
14.5595 | -105.523
14.5631 | -105.515
14.5668 | -105.5079
14.5704 | -105.5019
14.5741 | -105.4961
14.5778 | -105.4897
14.5814 | -105.4865
14.5851 | -105.4843
14.5888 | -105.481
14.5924 | -105.4785
14.5961 | -105.4772
14.5998 | -105.4773
14.6034 | -105.4761
14.6071 | -105.4755
14.6108 | -105.4766
14.6144 | -105.477
14.6181 | -105.4779
14.6218 | -105.4757
14.6254 | -105.4769
14.6291 | -105.4769
14.6328 | -105.4771
14.6364 | -105.475
14.6401 | -105.4747
14.6438 | -105.4768
14.6474 | -105.4788
14.6511 | -105.4857
14.6548 | -105.4905
14.6584 | -105.4955
14.6621 | -105.5046
14.6657 | -105.5113
14.6694 | -105.5156
14.6731 | -105.5209
14.6767 | -105.5258
14.6804 | -105.5322
14.6841 | -105.5385
14.6877 | -105.5433
14.6914 | -105.547
14.6951 | -105.5525
14.6987 | -105.5584
14.7024 | -105.5628
14.7061 | -105.5658
14.7097 | -105.5698
14.7134 | -105.5767
14.7171 | -105.5806
14.7207 | -105.5837
14.7244 | -105.5868
14.7281 | -105.587
14.7317 | -105.5839
14.7354 | -105.5827
14.7391 | -105.5856
14.7427 | -105.5897
14.7464 | -105.5941
14.7501 | -105.596
14.7537 | -105.5977
14.7574 | -105.5967
14.7611 | -105.5949
14.7647 | -105.5958
14.7684 | -105.6003
14.7721 | -105.6052
14.7757 | -105.6068
14.7794 | -105.6079
14.783 | -105.6071
14.7867 | -105.6067
14.7904 | -105.6069
14.794 | -105.6086
14.7977 | -105.6101
14.8014 | -105.6111
14.805 | -105.6089
14.8087 | -105.6106
14.8124 | -105.6108
14.816 | -105.6124
14.8197 | -105.6129
14.8234 | -105.6105
14.827 | -105.6107
14.8307 | -105.6126
14.8344 | -105.6177
14.838 | -105.6193
14.8417 | -105.6188
14.8454 | -105.616
14.849 | -105.6155
14.8527 | -105.6157
14.8564 | -105.6141
14.86 | -105.6135
14.8637 | -105.6128
14.8674 | -105.6127
14.871 | -105.6117
14.8747 | -105.6075
14.8784 | -105.6051
14.882 | -105.6043
14.8857 | -105.6047
14.8893 | -105.6036
14.893 | -105.6047
14.8967 | -105.6054
14.9003 | -105.6039
14.904 | -105.602
14.9077 | -105.5992
14.9113 | -105.5928
14.915 | -105.5882
14.9187 | -105.5849
14.9223 | -105.5788
14.926 | -105.5719
14.9297 | -105.5643
14.9333 | -105.5561
14.937 | -105.5412
14.9407 | -105.5254
14.9443 | -105.5143
14.948 | -105.5005
14.9517 | -105.4815
14.9553 | -105.4638
14.959 | -105.4448
14.9627 | -105.4243
14.9663 | -105.4057
14.97 | -105.3859
14.9737 | -105.3668
14.9773 | -105.3484
14.981 | -105.3258
14.9847 | -105.3054
14.9883 | -105.2857
14.992 | -105.265
14.9956 | -105.2471
14.9993 | -105.2291
15.003 | -105.2074
15.0066 | -105.1852
15.0103 | -105.1635
15.014 | -105.1437
15.0176 | -105.127
15.0213 | -105.1127
15.025 | -105.0976
15.0286 | -105.0825
15.0323 | -105.0671
15.036 | -105.0557
15.0396 | -105.0496
15.0433 | -105.0422
15.047 | -105.035
15.0506 | -105.0292
15.0543 | -105.0252
15.058 | -105.0255
15.0616 | -105.0283
15.0653 | -105.0307
15.069 | -105.0323
15.0726 | -105.035
15.0763 | -105.0388
15.08 | -105.044
15.0836 | -105.0513
15.0873 | -105.0607
15.091 | -105.0704
15.0946 | -105.079
15.0983 | -105.0876
15.1019 | -105.0955
15.1056 | -105.1048
15.1093 | -105.1129
15.1129 | -105.1202
15.1166 | -105.1285
15.1203 | -105.1362
15.1239 | -105.1422
15.1276 | -105.1494
15.1313 | -105.1559
15.1349 | -105.1612
15.1386 | -105.1683
15.1423 | -105.176
15.1459 | -105.1824
15.1496 | -105.1887
15.1533 | -105.1991
15.1569 | -105.2068
15.1606 | -105.2122
15.1643 | -105.2172
15.1679 | -105.2235
15.1716 | -105.2268
15.1753 | -105.2312
15.1789 | -105.2398
15.1826 | -105.2456
15.1863 | -105.2498
15.1899 | -105.254
15.1936 | -105.2574
15.1973 | -105.2613
15.2009 | -105.2675
15.2046 | -105.2749
15.2082 | -105.2819
15.2119 | -105.288
15.2156 | -105.295
15.2192 | -105.302
15.2229 | -105.3095
15.2266 | -105.3163
15.2302 | -105.3227
15.2339 | -105.3283
15.2376 | -105.3319
15.2412 | -105.3351
15.2449 | -105.341
15.2486 | -105.3444
15.2522 | -105.3444
15.2559 | -105.344
15.2596 | -105.3468
15.2632 | -105.3487
15.2669 | -105.3485
15.2706 | -105.3485
15.2742 | -105.3436
15.2779 | -105.3402
15.2816 | -105.3391
15.2852 | -105.3398
15.2889 | -105.338
15.2926 | -105.337
15.2962 | -105.3353
15.2999 | -105.3334
15.3036 | -105.3333
15.3072 | -105.3351
15.3109 | -105.3383
15.3145 | -105.3411
15.3182 | -105.3442
15.3219 | -105.3472
15.3255 | -105.3497
15.3292 | -105.3539
15.3329 | -105.3592
15.3365 | -105.3601
15.3402 | -105.3609
15.3439 | -105.3663
15.3475 | -105.3721
15.3512 | -105.376
15.3549 | -105.3797
15.3585 | -105.3826
15.3622 | -105.386
15.3659 | -105.3896
15.3695 | -105.3909
15.3732 | -105.3938
15.3769 | -105.3982
15.3805 | -105.4022
15.3842 | -105.4058
15.3879 | -105.4082
15.3915 | -105.4143
15.3952 | -105.4188
15.3989 | -105.4196
15.4025 | -105.4232
15.4062 | -105.4269
15.4099 | -105.4302
15.4135 | -105.4285
15.4172 | -105.4265
15.4208 | -105.4221
15.4245 | -105.4174
15.4282 | -105.4147
15.4318 | -105.4104
15.4355 | -105.4038
15.4392 | -105.3958
15.4428 | -105.386
15.4465 | -105.3764
15.4502 | -105.364
15.4538 | -105.3484
15.4575 | -105.334
15.4612 | -105.3166
15.4648 | -105.2977
15.4685 | -105.2815
15.4722 | -105.2626
15.4758 | -105.2404
15.4795 | -105.2174
15.4832 | -105.1962
15.4868 | -105.1755
15.4905 | -105.1567
15.4942 | -105.1396
15.4978 | -105.1205
15.5015 | -105.1015
15.5052 | -105.082
15.5088 | -105.0611
15.5125 | -105.0421
15.5162 | -105.0261
15.5198 | -105.0102
15.5235 | -104.9944
15.5271 | -104.9808
15.5308 | -104.9688
15.5345 | -104.9586
15.5381 | -104.9476
15.5418 | -104.938
15.5455 | -104.9293
15.5491 | -104.923
15.5528 | -104.9194
15.5565 | -104.9168
15.5601 | -104.9173
15.5638 | -104.9176
15.5675 | -104.9153
15.5711 | -104.9121
15.5748 | -104.9099
15.5785 | -104.9098
15.5821 | -104.9095
15.5858 | -104.9095
15.5895 | -104.9119
15.5931 | -104.9149
15.5968 | -104.9163
15.6005 | -104.9178
15.6041 | -104.9216
15.6078 | -104.9279
15.6115 | -104.9342
15.6151 | -104.9383
15.6188 | -104.9449
15.6225 | -104.9547
15.6261 | -104.9641
15.6298 | -104.9726
15.6334 | -104.9814
15.6371 | -104.9952
15.6408 | -105.0066
15.6444 | -105.0154
15.6481 | -105.0251
15.6518 | -105.0338
15.6554 | -105.0414
15.6591 | -105.0487
15.6628 | -105.0592
15.6664 | -105.0697
15.6701 | -105.0785
15.6738 | -105.0864
15.6774 | -105.0959
15.6811 | -105.1028
15.6848 | -105.1119
15.6884 | -105.1183
15.6921 | -105.1244
15.6958 | -105.1333
15.6994 | -105.1417
15.7031 | -105.1482
15.7068 | -105.1502
15.7104 | -105.1553
15.7141 | -105.1614
15.7178 | -105.1686
15.7214 | -105.1741
15.7251 | -105.1777
15.7288 | -105.1817
15.7324 | -105.1918
15.7361 | -105.1982
15.7397 | -105.2029
15.7434 | -105.2078
15.7471 | -105.2143
15.7507 | -105.2178
15.7544 | -105.2203
15.7581 | -105.2227
15.7617 | -105.2269
15.7654 | -105.2317
15.7691 | -105.2363
15.7727 | -105.236
15.7764 | -105.2353
15.7801 | -105.2355
15.7837 | -105.2366
15.7874 | -105.2402
15.7911 | -105.2447
15.7947 | -105.2482
15.7984 | -105.2516
15.8021 | -105.2556
15.8057 | -105.2607
15.8094 | -105.267
15.8131 | -105.2721
15.8167 | -105.2771
15.8204 | -105.282
15.8241 | -105.2858
15.8277 | -105.2866
15.8314 | -105.2881
15.8351 | -105.2923
15.8387 | -105.2976
15.8424 | -105.3026
15.846 | -105.3055
15.8497 | -105.3073
15.8534 | -105.3103
15.857 | -105.3146
15.8607 | -105.3156
15.8644 | -105.3164
15.868 | -105.3197
15.8717 | -105.3207
15.8754 | -105.3203
15.879 | -105.3178
15.8827 | -105.3158
15.8864 | -105.316
15.89 | -105.315
15.8937 | -105.3116
15.8974 | -105.3046
15.901 | -105.3003
15.9047 | -105.2961
15.9084 | -105.2915
15.912 | -105.2837
15.9157 | -105.2779
15.9194 | -105.2711
15.923 | -105.262
15.9267 | -105.2509
15.9304 | -105.2432
15.934 | -105.238
15.9377 | -105.2296
15.9414 | -105.2204
15.945 | -105.2115
15.9487 | -105.2036
15.9523 | -105.1961
15.956 | -105.1849
15.9597 | -105.1716
15.9633 | -105.1617
15.967 | -105.1529
15.9707 | -105.1451
15.9743 | -105.1344
15.978 | -105.1241
15.9817 | -105.1148
15.9853 | -105.1032
15.989 | -105.0935
15.9927 | -105.086
15.9963 | -105.0772
16 | -105.0715
16.0037 | -105.0647
16.0073 | -105.0532
16.011 | -105.0425
16.0147 | -105.0335
16.0183 | -105.0249
16.022 | -105.0178
16.0257 | -105.0121
16.0293 | -105.0043
16.033 | -104.9966
16.0367 | -104.988
16.0403 | -104.9793
16.044 | -104.9725
16.0477 | -104.9647
16.0513 | -104.9577
16.055 | -104.9529
16.0586 | -104.9446
16.0623 | -104.9336
16.066 | -104.9229
16.0696 | -104.9129
16.0733 | -104.9045
16.077 | -104.8968
16.0806 | -104.8902
16.0843 | -104.8803
16.088 | -104.8705
16.0916 | -104.8599
16.0953 | -104.8504
16.099 | -104.8369
16.1026 | -104.8283
16.1063 | -104.8222
16.11 | -104.8172
16.1136 | -104.8127
16.1173 | -104.8043
16.121 | -104.7972
16.1246 | -104.7934
16.1283 | -104.7932
16.132 | -104.7916
16.1356 | -104.7903
16.1393 | -104.7904
16.143 | -104.7915
16.1466 | -104.794
16.1503 | -104.7988
16.154 | -104.8037
16.1576 | -104.8109
16.1613 | -104.8163
16.1649 | -104.8219
16.1686 | -104.8296
16.1723 | -104.8376
16.1759 | -104.8471
16.1796 | -104.8589
16.1833 | -104.871
16.1869 | -104.8861
16.1906 | -104.9012
16.1943 | -104.9135
16.1979 | -104.9268
16.2016 | -104.9416
16.2053 | -104.9569
16.2089 | -104.9734
16.2126 | -104.9883
16.2163 | -105.0039
16.2199 | -105.0177
16.2236 | -105.0274
16.2273 | -105.0373
16.2309 | -105.0459
16.2346 | -105.0577
16.2383 | -105.0715
16.2419 | -105.0828
16.2456 | -105.091
16.2493 | -105.0961
16.2529 | -105.1008
16.2566 | -105.1052
16.2603 | -105.1078
16.2639 | -105.1105
16.2676 | -105.1141
16.2712 | -105.1185
16.2749 | -105.1195
16.2786 | -105.1209
16.2822 | -105.1203
16.2859 | -105.1184
16.2896 | -105.119
16.2932 | -105.1201
16.2969 | -105.1185
16.3006 | -105.1161
16.3042 | -105.1128
16.3079 | -105.1082
16.3116 | -105.1033
16.3152 | -105.1021
16.3189 | -105.1015
16.3226 | -105.1019
16.3262 | -105.1026
16.3299 | -105.1051
16.3336 | -105.1047
16.3372 | -105.1033
16.3409 | -105.1036
16.3446 | -105.106
16.3482 | -105.1104
16.3519 | -105.1149
16.3556 | -105.1222
16.3592 | -105.1303
16.3629 | -105.1383
16.3666 | -105.1465
16.3702 | -105.1591
16.3739 | -105.1735
16.3775 | -105.1877
16.3812 | -105.2007
16.3849 | -105.2146
16.3885 | -105.229
16.3922 | -105.2423
16.3959 | -105.2533
16.3995 | -105.2623
16.4032 | -105.2711
16.4069 | -105.2805
16.4105 | -105.2871
16.4142 | -105.2909
16.4179 | -105.2932
16.4215 | -105.2942
16.4252 | -105.2925
16.4289 | -105.291
16.4325 | -105.2862
16.4362 | -105.2797
16.4399 | -105.2718
16.4435 | -105.262
16.4472 | -105.2532
16.4509 | -105.2465
16.4545 | -105.2338
16.4582 | -105.2213
16.4619 | -105.2096
16.4655 | -105.1974
16.4692 | -105.1813
16.4729 | -105.166
16.4765 | -105.1537
16.4802 | -105.1429
16.4838 | -105.1322
16.4875 | -105.1211
16.4912 | -105.1087
16.4948 | -105.0994
16.4985 | -105.0888
16.5022 | -105.0785
16.5058 | -105.0703
16.5095 | -105.0656
16.5132 | -105.0605
16.5168 | -105.058
16.5205 | -105.0598
16.5242 | -105.062
16.5278 | -105.0669
16.5315 | -105.0762
16.5352 | -105.0869
16.5388 | -105.0997
16.5425 | -105.1168
16.5462 | -105.135
16.5498 | -105.1515
16.5535 | -105.1717
16.5572 | -105.1916
16.5608 | -105.2133
16.5645 | -105.2388
16.5682 | -105.2674
16.5718 | -105.2941
16.5755 | -105.3193
16.5792 | -105.346
16.5828 | -105.3751
16.5865 | -105.4023
16.5901 | -105.4257
16.5938 | -105.4503
16.5975 | -105.4783
16.6011 | -105.5072
16.6048 | -105.5332
16.6085 | -105.5585
16.6121 | -105.581
16.6158 | -105.6019
16.6195 | -105.6223
16.6231 | -105.6414
16.6268 | -105.6596
16.6305 | -105.6765
16.6341 | -105.694
16.6378 | -105.7114
16.6415 | -105.7242
16.6451 | -105.7323
16.6488 | -105.7395
16.6525 | -105.7474
16.6561 | -105.7511
16.6598 | -105.7535
16.6635 | -105.7557
16.6671 | -105.755
16.6708 | -105.754
16.6745 | -105.7514
16.6781 | -105.7475
16.6818 | -105.7407
16.6855 | -105.7297
16.6891 | -105.7176
16.6928 | -105.7106
16.6964 | -105.7018
16.7001 | -105.6925
16.7038 | -105.6832
16.7074 | -105.6722
16.7111 | -105.6597
16.7148 | -105.6446
16.7184 | -105.6289
16.7221 | -105.6178
16.7258 | -105.6088
16.7294 | -105.6008
16.7331 | -105.5931
16.7368 | -105.5857
16.7404 | -105.5794
16.7441 | -105.5731
16.7478 | -105.5668
16.7514 | -105.5629
16.7551 | -105.5592
16.7588 | -105.5539
16.7624 | -105.5516
16.7661 | -105.5495
16.7698 | -105.5472
16.7734 | -105.5471
16.7771 | -105.5477
16.7808 | -105.5459
16.7844 | -105.5445
16.7881 | -105.5494
16.7918 | -105.5554
16.7954 | -105.5618
16.7991 | -105.5666
16.8027 | -105.5689
16.8064 | -105.5701
16.8101 | -105.5739
16.8137 | -105.5786
16.8174 | -105.5791
16.8211 | -105.5792
16.8247 | -105.5811
16.8284 | -105.5832
16.8321 | -105.5825
16.8357 | -105.5795
16.8394 | -105.5747
16.8431 | -105.5694
16.8467 | -105.5647
16.8504 | -105.5599
16.8541 | -105.5529
16.8577 | -105.5452
16.8614 | -105.5343
16.8651 | -105.521
16.8687 | -105.5108
16.8724 | -105.5017
16.8761 | -105.492
16.8797 | -105.4822
16.8834 | -105.4734
16.8871 | -105.4632
16.8907 | -105.4514
16.8944 | -105.4418
16.8981 | -105.4333
16.9017 | -105.4235
16.9054 | -105.4132
16.909 | -105.4029
16.9127 | -105.3914
16.9164 | -105.3823
16.92 | -105.3789
16.9237 | -105.3717
16.9274 | -105.3679
16.931 | -105.3645
16.9347 | -105.3601
16.9384 | -105.3579
16.942 | -105.3571
16.9457 | -105.3568
16.9494 | -105.3601
16.953 | -105.3632
16.9567 | -105.3674
16.9604 | -105.3749
16.964 | -105.3812
16.9677 | -105.3881
16.9714 | -105.3948
16.975 | -105.4015
16.9787 | -105.4128
16.9824 | -105.4235
16.986 | -105.4354
16.9897 | -105.4491
16.9934 | -105.4645
16.997 | -105.4787
17.0007 | -105.4909
17.0044 | -105.5043
17.008 | -105.5191
17.0117 | -105.5326
17.0153 | -105.5461
17.019 | -105.5584
17.0227 | -105.5732
17.0263 | -105.587
17.03 | -105.5962
17.0337 | -105.6045
17.0373 | -105.6129
17.041 | -105.6234
17.0447 | -105.6327
17.0483 | -105.6388
17.052 | -105.642
17.0557 | -105.6439
17.0593 | -105.6469
17.063 | -105.6454
17.0667 | -105.6399
17.0703 | -105.6362
17.074 | -105.6316
17.0777 | -105.6248
17.0813 | -105.6163
17.085 | -105.6047
17.0887 | -105.5946
17.0923 | -105.5836
17.096 | -105.5725
17.0997 | -105.5607
17.1033 | -105.5486
17.107 | -105.5345
17.1107 | -105.5175
17.1143 | -105.4978
17.118 | -105.478
17.1216 | -105.4589
17.1253 | -105.4386
17.129 | -105.4215
17.1326 | -105.4042
17.1363 | -105.3871
17.14 | -105.3688
17.1436 | -105.351
17.1473 | -105.3328
17.151 | -105.3175
17.1546 | -105.3046
17.1583 | -105.2917
17.162 | -105.2792
17.1656 | -105.2689
17.1693 | -105.2599
17.173 | -105.2503
17.1766 | -105.2398
17.1803 | -105.231
17.184 | -105.2232
17.1876 | -105.2184
17.1913 | -105.2151
17.195 | -105.2122
17.1986 | -105.2125
17.2023 | -105.2145
17.206 | -105.2155
17.2096 | -105.2173
17.2133 | -105.2213
17.217 | -105.2287
17.2206 | -105.237
17.2243 | -105.2427
17.2279 | -105.2518
17.2316 | -105.2608
17.2353 | -105.2655
17.2389 | -105.272
17.2426 | -105.2814
17.2463 | -105.2925
17.2499 | -105.3
17.2536 | -105.3068
17.2573 | -105.3146
17.2609 | -105.3182
17.2646 | -105.3213
17.2683 | -105.3209
17.2719 | -105.3216
17.2756 | -105.3207
17.2793 | -105.3147
17.2829 | -105.3072
17.2866 | -105.298
17.2903 | -105.2895
17.2939 | -105.2794
17.2976 | -105.2671
17.3013 | -105.2543
17.3049 | -105.2391
17.3086 | -105.2264
17.3123 | -105.2142
17.3159 | -105.1963
17.3196 | -105.1775
17.3233 | -105.1583
17.3269 | -105.1405
17.3306 | -105.1243
17.3343 | -105.1058
17.3379 | -105.0878
17.3416 | -105.0699
17.3452 | -105.0549
17.3489 | -105.0401
17.3526 | -105.0283
17.3562 | -105.0179
17.3599 | -105.0085
17.3636 | -105
17.3672 | -104.995
17.3709 | -104.9907
17.3746 | -104.9847
17.3782 | -104.9808
17.3819 | -104.9787
17.3856 | -104.9788
17.3892 | -104.982
17.3929 | -104.9876
17.3966 | -104.9926
17.4002 | -104.9994
17.4039 | -105.0081
17.4076 | -105.0172
17.4112 | -105.025
17.4149 | -105.0348
17.4186 | -105.0459
17.4222 | -105.0589
17.4259 | -105.0733
17.4296 | -105.0871
17.4332 | -105.0996
17.4369 | -105.1116
17.4406 | -105.1246
17.4442 | -105.138
17.4479 | -105.1476
17.4515 | -105.16
17.4552 | -105.1723
17.4589 | -105.1852
17.4625 | -105.1945
17.4662 | -105.2025
17.4699 | -105.2105
17.4735 | -105.2197
17.4772 | -105.2277
17.4809 | -105.234
17.4845 | -105.2428
17.4882 | -105.247
17.4919 | -105.2492
17.4955 | -105.2496
17.4992 | -105.2498
17.5029 | -105.25
17.5065 | -105.2474
17.5102 | -105.2449
17.5139 | -105.243
17.5175 | -105.2397
17.5212 | -105.2327
17.5249 | -105.2242
17.5285 | -105.2181
17.5322 | -105.2108
17.5359 | -105.2032
17.5395 | -105.193
17.5432 | -105.1817
17.5468 | -105.1731
17.5505 | -105.1648
17.5542 | -105.1561
17.5578 | -105.1456
17.5615 | -105.1358
17.5652 | -105.1271
17.5688 | -105.119
17.5725 | -105.1121
17.5762 | -105.1073
17.5798 | -105.1027
17.5835 | -105.0964
17.5872 | -105.0916
17.5908 | -105.09
17.5945 | -105.0892
17.5982 | -105.0908
17.6018 | -105.0933
17.6055 | -105.0961
17.6092 | -105.0997
17.6128 | -105.1067
17.6165 | -105.113
17.6202 | -105.1212
17.6238 | -105.1332
17.6275 | -105.1465
17.6312 | -105.1598
17.6348 | -105.173
17.6385 | -105.1873
17.6422 | -105.2028
17.6458 | -105.2179
17.6495 | -105.2342
17.6532 | -105.2512
17.6568 | -105.2682
17.6605 | -105.2872
17.6641 | -105.306
17.6678 | -105.3229
17.6715 | -105.3393
17.6751 | -105.3573
17.6788 | -105.3732
17.6825 | -105.3874
17.6861 | -105.403
17.6898 | -105.4156
17.6935 | -105.4249
17.6971 | -105.4348
17.7008 | -105.4428
17.7045 | -105.4489
17.7081 | -105.4525
17.7118 | -105.4516
17.7155 | -105.4508
17.7191 | -105.4506
17.7228 | -105.4489
17.7265 | -105.4445
17.7301 | -105.4378
17.7338 | -105.4308
17.7375 | -105.4232
17.7411 | -105.4147
17.7448 | -105.4052
17.7485 | -105.396
17.7521 | -105.3884
17.7558 | -105.376
17.7595 | -105.366
17.7631 | -105.356
17.7668 | -105.3442
17.7704 | -105.3328
17.7741 | -105.3244
17.7778 | -105.3161
17.7814 | -105.3078
17.7851 | -105.2995
17.7888 | -105.2911
17.7924 | -105.2842
17.7961 | -105.2783
17.7998 | -105.2725
17.8034 | -105.2679
17.8071 | -105.2633
17.8108 | -105.2601
17.8144 | -105.2567
17.8181 | -105.2564
17.8218 | -105.2564
17.8254 | -105.2544
17.8291 | -105.2527
17.8328 | -105.2508
17.8364 | -105.2525
17.8401 | -105.2555
17.8438 | -105.2551
17.8474 | -105.2502
17.8511 | -105.2499
17.8548 | -105.2522
17.8584 | -105.2548
17.8621 | -105.2547
17.8657 | -105.2513
17.8694 | -105.2492
17.8731 | -105.2472
17.8767 | -105.2417
17.8804 | -105.2364
17.8841 | -105.2337
17.8877 | -105.2299
17.8914 | -105.2239
17.8951 | -105.2161
17.8987 | -105.2115
17.9024 | -105.205
17.9061 | -105.1961
17.9097 | -105.1869
17.9134 | -105.1761
17.9171 | -105.1634
17.9207 | -105.1519
17.9244 | -105.1389
17.9281 | -105.1219
17.9317 | -105.109
17.9354 | -105.0926
17.9391 | -105.0744
17.9427 | -105.058
17.9464 | -105.044
17.9501 | -105.0296
17.9537 | -105.0128
17.9574 | -104.9971
17.9611 | -104.9815
17.9647 | -104.9658
17.9684 | -104.9514
17.972 | -104.9352
17.9757 | -104.9195
17.9794 | -104.9058
17.983 | -104.8911
17.9867 | -104.8751
17.9904 | -104.8606
17.994 | -104.8469
17.9977 | -104.8343
18.0014 | -104.8217
18.005 | -104.8089
18.0087 | -104.7965
18.0124 | -104.7877
18.016 | -104.7777
18.0197 | -104.7689
18.0234 | -104.7585
18.027 | -104.747
18.0307 | -104.735
18.0344 | -104.722
18.038 | -104.7084
18.0417 | -104.6986
18.0454 | -104.6895
18.049 | -104.6808
18.0527 | -104.6737
18.0564 | -104.6669
18.06 | -104.6587
18.0637 | -104.6515
18.0674 | -104.6466
18.071 | -104.642
18.0747 | -104.6383
18.0784 | -104.6349
18.082 | -104.632
18.0857 | -104.6286
18.0893 | -104.6233
18.093 | -104.6169
18.0967 | -104.6099
18.1003 | -104.6021
18.104 | -104.5965
18.1077 | -104.5856
18.1113 | -104.5724
18.115 | -104.5585
18.1187 | -104.5481
18.1223 | -104.5375
18.126 | -104.5249
18.1297 | -104.5143
18.1333 | -104.5043
18.137 | -104.4908
18.1407 | -104.474
18.1443 | -104.4531
18.148 | -104.4333
18.1517 | -104.4162
18.1553 | -104.3948
18.159 | -104.3746
18.1627 | -104.3546
18.1663 | -104.3357
18.17 | -104.3189
18.1737 | -104.3026
18.1773 | -104.2847
18.181 | -104.2668
18.1847 | -104.2509
18.1883 | -104.2336
18.192 | -104.2216
18.1956 | -104.2106
18.1993 | -104.1988
18.203 | -104.1863
18.2066 | -104.1758
18.2103 | -104.1686
18.214 | -104.163
18.2176 | -104.1584
18.2213 | -104.156
18.225 | -104.1528
18.2286 | -104.1506
18.2323 | -104.1502
18.236 | -104.1523
18.2396 | -104.1556
18.2433 | -104.1601
18.247 | -104.1668
18.2506 | -104.1719
18.2543 | -104.1791
18.258 | -104.1877
18.2616 | -104.197
18.2653 | -104.2054
18.269 | -104.214
18.2726 | -104.2212
18.2763 | -104.2285
18.28 | -104.236
18.2836 | -104.246
18.2873 | -104.2556
18.291 | -104.2635
18.2946 | -104.2737
18.2983 | -104.2829
18.3019 | -104.2891
18.3056 | -104.2937
18.3093 | -104.2983
18.3129 | -104.3045
18.3166 | -104.3099
18.3203 | -104.3165
18.3239 | -104.3223
18.3276 | -104.3275
18.3313 | -104.335
18.3349 | -104.3417
18.3386 | -104.3496
18.3423 | -104.3603
18.3459 | -104.3719
18.3496 | -104.3849
18.3533 | -104.3985
18.3569 | -104.414
18.3606 | -104.4321
18.3643 | -104.447
18.3679 | -104.4628
18.3716 | -104.4812
18.3753 | -104.5009
18.3789 | -104.5184
18.3826 | -104.5348
18.3863 | -104.5498
18.3899 | -104.5626
18.3936 | -104.5748
18.3973 | -104.5877
18.4009 | -104.6031
18.4046 | -104.6209
18.4082 | -104.6385
18.4119 | -104.6503
18.4156 | -104.6626
18.4192 | -104.675
18.4229 | -104.685
18.4266 | -104.6923
18.4302 | -104.6985
18.4339 | -104.7049
18.4376 | -104.7107
18.4412 | -104.7162
18.4449 | -104.7208
18.4486 | -104.7249
18.4522 | -104.7271
18.4559 | -104.7284
18.4596 | -104.7327
18.4632 | -104.738
18.4669 | -104.7414
18.4706 | -104.7439
18.4742 | -104.7462
18.4779 | -104.7508
18.4816 | -104.7568
18.4852 | -104.7629
18.4889 | -104.7673
18.4926 | -104.7724
18.4962 | -104.778
18.4999 | -104.7841
18.5036 | -104.7886
18.5072 | -104.7965
18.5109 | -104.8027
18.5145 | -104.8111
18.5182 | -104.8176
18.5219 | -104.8232
18.5255 | -104.8295
18.5292 | -104.8361
18.5329 | -104.8423
18.5365 | -104.8513
18.5402 | -104.8562
18.5439 | -104.861
18.5475 | -104.8661
18.5512 | -104.8714
18.5549 | -104.8752
18.5585 | -104.8771
18.5622 | -104.878
18.5659 | -104.8793
18.5695 | -104.8822
18.5732 | -104.8836
18.5769 | -104.8821
18.5805 | -104.881
18.5842 | -104.8786
18.5879 | -104.8764
18.5915 | -104.8715
18.5952 | -104.8656
18.5989 | -104.8613
18.6025 | -104.8554
18.6062 | -104.8475
18.6099 | -104.8387
18.6135 | -104.8303
18.6172 | -104.8205
18.6208 | -104.8094
18.6245 | -104.7977
18.6282 | -104.7867
18.6318 | -104.7764
18.6355 | -104.7639
18.6392 | -104.7524
18.6428 | -104.7424
18.6465 | -104.7325
18.6502 | -104.7235
18.6538 | -104.7155
18.6575 | -104.7094
18.6612 | -104.7022
18.6648 | -104.6949
18.6685 | -104.6889
18.6722 | -104.6822
18.6758 | -104.679
18.6795 | -104.6758
18.6832 | -104.6724
18.6868 | -104.6719
18.6905 | -104.6736
18.6942 | -104.6744
18.6978 | -104.6755
18.7015 | -104.6794
18.7052 | -104.684
18.7088 | -104.6893
18.7125 | -104.695
18.7162 | -104.7004
18.7198 | -104.7083
18.7235 | -104.7155
18.7271 | -104.7231
18.7308 | -104.7326
18.7345 | -104.7448
18.7381 | -104.7563
18.7418 | -104.7643
18.7455 | -104.7723
18.7491 | -104.7832
18.7528 | -104.7966
18.7565 | -104.8087
18.7601 | -104.8191
18.7638 | -104.8298
18.7675 | -104.8419
18.7711 | -104.8521
18.7748 | -104.86
18.7785 | -104.8674
18.7821 | -104.8739
18.7858 | -104.881
18.7895 | -104.8878
18.7931 | -104.8952
18.7968 | -104.9015
18.8005 | -104.9057
18.8041 | -104.9084
18.8078 | -104.914
18.8115 | -104.9176
18.8151 | -104.9183
18.8188 | -104.92
18.8225 | -104.9217
18.8261 | -104.9207
18.8298 | -104.9206
18.8334 | -104.9183
18.8371 | -104.9142
18.8408 | -104.9102
18.8444 | -104.9072
18.8481 | -104.9036
18.8518 | -104.9014
18.8554 | -104.9008
18.8591 | -104.9005
18.8628 | -104.9
18.8664 | -104.8976
18.8701 | -104.8961
18.8738 | -104.8979
18.8774 | -104.8996
18.8811 | -104.9032
18.8848 | -104.9084
18.8884 | -104.912
18.8921 | -104.9173
18.8958 | -104.9252
18.8994 | -104.932
18.9031 | -104.9392
18.9068 | -104.9465
18.9104 | -104.9536
18.9141 | -104.9633
18.9178 | -104.9751
18.9214 | -104.9859
18.9251 | -104.9968
18.9288 | -105.0107
18.9324 | -105.0231
18.9361 | -105.0348
18.9397 | -105.0487
18.9434 | -105.06
18.9471 | -105.0716
18.9507 | -105.0845
18.9544 | -105.0979
18.9581 | -105.1088
18.9617 | -105.1219
18.9654 | -105.1343
18.9691 | -105.1446
18.9727 | -105.1568
18.9764 | -105.1686
18.9801 | -105.1804
18.9837 | -105.1916
18.9874 | -105.2003
18.9911 | -105.2089
18.9947 | -105.2171
18.9984 | -105.2232
19.0021 | -105.2254
19.0057 | -105.2303
19.0094 | -105.2382
19.0131 | -105.2442
19.0167 | -105.2434
19.0204 | -105.2415
19.0241 | -105.2405
19.0277 | -105.2384
19.0314 | -105.2362
19.0351 | -105.2283
19.0387 | -105.2256
19.0424 | -105.2223
19.046 | -105.2207
19.0497 | -105.2156
19.0534 | -105.2087
19.057 | -105.2012
19.0607 | -105.1913
19.0644 | -105.1808
19.068 | -105.1722
19.0717 | -105.1632
19.0754 | -105.1521
19.079 | -105.1405
19.0827 | -105.1291
19.0864 | -105.118
19.09 | -105.1053
19.0937 | -105.0918
19.0974 | -105.0795
19.101 | -105.0632
19.1047 | -105.049
19.1084 | -105.0349
19.112 | -105.0221
19.1157 | -105.0097
19.1194 | -104.9956
19.123 | -104.9819
19.1267 | -104.9689
19.1304 | -104.9573
19.134 | -104.9474
19.1377 | -104.9378
19.1414 | -104.9279
19.145 | -104.9172
19.1487 | -104.9093
19.1523 | -104.9018
19.156 | -104.8955
19.1597 | -104.8916
19.1633 | -104.8882
19.167 | -104.8838
19.1707 | -104.8791
19.1743 | -104.8753
19.178 | -104.8735
19.1817 | -104.8723
19.1853 | -104.8717
19.189 | -104.8698
19.1927 | -104.8666
19.1963 | -104.8637
19.2 | -104.8611
19.2037 | -104.8612
19.2073 | -104.8611
19.211 | -104.8594
19.2147 | -104.8606
19.2183 | -104.86
19.222 | -104.8568
19.2257 | -104.8526
19.2293 | -104.848
19.233 | -104.8418
19.2367 | -104.8365
19.2403 | -104.8323
19.244 | -104.8258
19.2477 | -104.8175
19.2513 | -104.8092
19.255 | -104.8024
19.2586 | -104.7964
19.2623 | -104.7886
19.266 | -104.7807
19.2696 | -104.7747
19.2733 | -104.7695
19.277 | -104.7645
19.2806 | -104.7597
19.2843 | -104.751
19.288 | -104.7466
19.2916 | -104.7434
19.2953 | -104.7377
19.299 | -104.7326
19.3026 | -104.7288
19.3063 | -104.7233
19.31 | -104.7183
19.3136 | -104.7134
19.3173 | -104.7089
19.321 | -104.7037
19.3246 | -104.7015
19.3283 | -104.7006
19.332 | -104.7002
19.3356 | -104.7011
19.3393 | -104.702
19.343 | -104.7032
19.3466 | -104.7057
19.3503 | -104.7082
19.354 | -104.7107
19.3576 | -104.7126
19.3613 | -104.7148
19.3649 | -104.7203
19.3686 | -104.7258
19.3723 | -104.7306
19.3759 | -104.7361
19.3796 | -104.7426
19.3833 | -104.7474
19.3869 | -104.7498
19.3906 | -104.752
19.3943 | -104.7572
19.3979 | -104.7623
19.4016 | -104.7664
19.4053 | -104.7687
19.4089 | -104.7688
19.4126 | -104.7722
19.4163 | -104.7758
19.4199 | -104.776
19.4236 | -104.7744
19.4273 | -104.7733
19.4309 | -104.7716
19.4346 | -104.7699
19.4383 | -104.7674
19.4419 | -104.7664
19.4456 | -104.7671
19.4493 | -104.766
19.4529 | -104.7623
19.4566 | -104.7584
19.4603 | -104.7533
19.4639 | -104.7478
19.4676 | -104.7458
19.4712 | -104.744
19.4749 | -104.7415
19.4786 | -104.7363
19.4822 | -104.7318
19.4859 | -104.7276
19.4896 | -104.7233
19.4932 | -104.7217
19.4969 | -104.7227
19.5006 | -104.7235
19.5042 | -104.727
19.5079 | -104.7317
19.5116 | -104.7349
19.5152 | -104.7375
19.5189 | -104.7413
19.5226 | -104.7448
19.5262 | -104.7492
19.5299 | -104.7543
19.5336 | -104.7589
19.5372 | -104.765
19.5409 | -104.7721
19.5446 | -104.7771
19.5482 | -104.7842
19.5519 | -104.7892
19.5556 | -104.7945
19.5592 | -104.8045
19.5629 | -104.8115
19.5666 | -104.8204
19.5702 | -104.8263
19.5739 | -104.8338
19.5775 | -104.8418
19.5812 | -104.8476
19.5849 | -104.8542
19.5885 | -104.8605
19.5922 | -104.8666
19.5959 | -104.8731
19.5995 | -104.8783
19.6032 | -104.8806
19.6069 | -104.8828
19.6105 | -104.8825
19.6142 | -104.8816
19.6179 | -104.8824
19.6215 | -104.8829
19.6252 | -104.8812
19.6289 | -104.8798
19.6325 | -104.8771
19.6362 | -104.8732
19.6399 | -104.8693
19.6435 | -104.8629
19.6472 | -104.8574
19.6509 | -104.8531
19.6545 | -104.8483
19.6582 | -104.8414
19.6619 | -104.8353
19.6655 | -104.8306
19.6692 | -104.824
19.6729 | -104.8158
19.6765 | -104.808
19.6802 | -104.7997
19.6838 | -104.7943
19.6875 | -104.7894
19.6912 | -104.7806
19.6948 | -104.7735
19.6985 | -104.7691
19.7022 | -104.7656
19.7058 | -104.7623
19.7095 | -104.7578
19.7132 | -104.7522
19.7168 | -104.7471
19.7205 | -104.7434
19.7242 | -104.7388
19.7278 | -104.7328
19.7315 | -104.7281
19.7352 | -104.725
19.7388 | -104.7221
19.7425 | -104.717
19.7462 | -104.7108
19.7498 | -104.7043
19.7535 | -104.6978
19.7572 | -104.6916
19.7608 | -104.6861
19.7645 | -104.682
19.7682 | -104.6789
19.7718 | -104.6739
19.7755 | -104.6653
19.7792 | -104.66
19.7828 | -104.655
19.7865 | -104.6479
19.7901 | -104.6424
19.7938 | -104.6374
19.7975 | -104.6329
19.8011 | -104.6309
19.8048 | -104.6281
19.8085 | -104.6245
19.8121 | -104.6229
19.8158 | -104.6223
19.8195 | -104.6197
19.8231 | -104.6171
19.8268 | -104.6153
19.8305 | -104.6148
19.8341 | -104.6132
19.8378 | -104.6118
19.8415 | -104.6116
19.8451 | -104.6113
19.8488 | -104.6106
19.8525 | -104.6114
19.8561 | -104.6123
19.8598 | -104.612
19.8635 | -104.6112
19.8671 | -104.6138
19.8708 | -104.6165
19.8745 | -104.6208
19.8781 | -104.6259
19.8818 | -104.6309
19.8855 | -104.6387
19.8891 | -104.6449
19.8928 | -104.651
19.8964 | -104.6577
19.9001 | -104.6624
19.9038 | -104.6656
19.9074 | -104.6724
19.9111 | -104.6811
19.9148 | -104.6903
19.9184 | -104.6986
19.9221 | -104.7068
19.9258 | -104.7152
19.9294 | -104.7243
19.9331 | -104.732
19.9368 | -104.7383
19.9404 | -104.7425
19.9441 | -104.7495
19.9478 | -104.7578
19.9514 | -104.766
19.9551 | -104.7714
19.9588 | -104.7762
19.9624 | -104.7813
19.9661 | -104.7854
19.9698 | -104.7885
19.9734 | -104.7889
19.9771 | -104.7884
19.9808 | -104.7876
19.9844 | -104.7847
19.9881 | -104.7809
19.9918 | -104.7772
19.9954 | -104.7717
19.9991 | -104.767
20.0027 | -104.7605
20.0064 | -104.7527
20.0101 | -104.7469
20.0137 | -104.7381
20.0174 | -104.7282
20.0211 | -104.719
20.0247 | -104.7111
20.0284 | -104.7023
20.0321 | -104.694
20.0357 | -104.6853
20.0394 | -104.6769
20.0431 | -104.6681
20.0467 | -104.6571
20.0504 | -104.6437
20.0541 | -104.6334
20.0577 | -104.6263
20.0614 | -104.616
20.0651 | -104.6064
20.0687 | -104.5979
20.0724 | -104.5887
20.0761 | -104.5819
20.0797 | -104.5748
20.0834 | -104.5675
20.0871 | -104.5623
20.0907 | -104.5551
20.0944 | -104.5519
20.0981 | -104.5468
20.1017 | -104.5429
20.1054 | -104.5422
20.109 | -104.5438
20.1127 | -104.5417
20.1164 | -104.5392
20.12 | -104.539
20.1237 | -104.5384
20.1274 | -104.5386
20.131 | -104.539
20.1347 | -104.5387
20.1384 | -104.5393
20.142 | -104.5411
20.1457 | -104.5444
20.1494 | -104.5495
20.153 | -104.5521
20.1567 | -104.5547
20.1604 | -104.5552
20.164 | -104.555
20.1677 | -104.5557
20.1714 | -104.5564
20.175 | -104.5601
20.1787 | -104.5647
20.1824 | -104.5688
20.186 | -104.5714
20.1897 | -104.5756
20.1934 | -104.5772
20.197 | -104.5783
20.2007 | -104.5813
20.2044 | -104.5849
20.208 | -104.5886
20.2117 | -104.59
20.2153 | -104.5908
20.219 | -104.5916
20.2227 | -104.5942
20.2263 | -104.5971
20.23 | -104.5969
20.2337 | -104.593
20.2373 | -104.591
20.241 | -104.5917
20.2447 | -104.5906
20.2483 | -104.5891
20.252 | -104.5855
20.2557 | -104.5814
20.2593 | -104.5793
20.263 | -104.5773
20.2667 | -104.5739
20.2703 | -104.5695
20.274 | -104.5638
20.2777 | -104.5567
20.2813 | -104.5507
20.285 | -104.5433
20.2887 | -104.5344
20.2923 | -104.5267
20.296 | -104.5166
20.2997 | -104.5058
20.3033 | -104.4922
20.307 | -104.4766
20.3107 | -104.4621
20.3143 | -104.4465
20.318 | -104.4298
20.3216 | -104.4144
20.3253 | -104.3987
20.329 | -104.3815
20.3326 | -104.3652
20.3363 | -104.351
20.34 | -104.3356
20.3436 | -104.3193
20.3473 | -104.3049
20.351 | -104.2914
20.3546 | -104.2785
20.3583 | -104.2661
20.362 | -104.2538
20.3656 | -104.243
20.3693 | -104.2322
20.373 | -104.2214
20.3766 | -104.2123
20.3803 | -104.2023
20.384 | -104.1924
20.3876 | -104.1836
20.3913 | -104.1743
20.395 | -104.1656
20.3986 | -104.1598
20.4023 | -104.1569
20.406 | -104.1573
20.4096 | -104.1571
20.4133 | -104.1565
20.417 | -104.1564
20.4206 | -104.1583
20.4243 | -104.163
20.4279 | -104.1684
20.4316 | -104.1728
20.4353 | -104.1765
20.4389 | -104.1807
20.4426 | -104.1864
20.4463 | -104.1904
20.4499 | -104.1949
20.4536 | -104.2016
20.4573 | -104.2093
20.4609 | -104.2163
20.4646 | -104.222
20.4683 | -104.2292
20.4719 | -104.2378
20.4756 | -104.2464
20.4793 | -104.2534
20.4829 | -104.2585
20.4866 | -104.2628
20.4903 | -104.267
20.4939 | -104.2702
20.4976 | -104.272
20.5013 | -104.2738
20.5049 | -104.2749
20.5086 | -104.2769
20.5123 | -104.2791
20.5159 | -104.2807
20.5196 | -104.2823
20.5233 | -104.284
20.5269 | -104.2831
20.5306 | -104.2819
20.5343 | -104.2809
20.5379 | -104.2768
20.5416 | -104.2727
20.5452 | -104.2695
20.5489 | -104.2642
20.5526 | -104.2587
20.5562 | -104.2528
20.5599 | -104.2428
20.5636 | -104.2339
20.5672 | -104.2266
20.5709 | -104.2181
20.5746 | -104.2094
20.5782 | -104.2022
20.5819 | -104.1936
20.5856 | -104.1842
20.5892 | -104.1757
20.5929 | -104.168
20.5966 | -104.161
20.6002 | -104.1525
20.6039 | -104.1412
20.6076 | -104.1294
20.6112 | -104.1182
20.6149 | -104.1081
20.6186 | -104.0978
20.6222 | -104.0901
20.6259 | -104.0799
20.6296 | -104.0683
20.6332 | -104.0563
20.6369 | -104.0458
20.6406 | -104.0358
20.6442 | -104.0246
20.6479 | -104.0153
20.6515 | -104.0061
20.6552 | -103.9972
20.6589 | -103.9884
20.6625 | -103.9803
20.6662 | -103.9723
20.6699 | -103.9664
20.6735 | -103.9619
20.6772 | -103.9566
20.6809 | -103.9517
20.6845 | -103.9485
20.6882 | -103.9446
20.6919 | -103.9392
20.6955 | -103.9345
20.6992 | -103.9281
20.7029 | -103.9229
20.7065 | -103.9194
20.7102 | -103.915
20.7139 | -103.9089
20.7175 | -103.9027
20.7212 | -103.8976
20.7249 | -103.8934
20.7285 | -103.8883
20.7322 | -103.8822
20.7359 | -103.8766
20.7395 | -103.8734
20.7432 | -103.8707
20.7468 | -103.8659
20.7505 | -103.8578
20.7542 | -103.8496
20.7578 | -103.843
20.7615 | -103.8363
20.7652 | -103.8281
20.7688 | -103.8195
20.7725 | -103.8114
20.7762 | -103.8042
20.7798 | -103.7966
20.7835 | -103.7891
20.7872 | -103.7819
20.7908 | -103.7741
20.7945 | -103.7649
20.7982 | -103.7554
20.8018 | -103.7453
20.8055 | -103.7333
20.8092 | -103.7213
20.8128 | -103.7082
20.8165 | -103.6941
20.8202 | -103.6775
20.8238 | -103.6599
20.8275 | -103.6429
20.8312 | -103.6274
20.8348 | -103.6126
20.8385 | -103.5967
20.8422 | -103.5807
20.8458 | -103.5623
20.8495 | -103.5442
20.8531 | -103.5297
20.8568 | -103.5127
20.8605 | -103.4946
20.8641 | -103.4776
20.8678 | -103.4595
20.8715 | -103.4416
20.8751 | -103.4268
20.8788 | -103.4134
20.8825 | -103.3988
20.8861 | -103.3852
20.8898 | -103.3727
20.8935 | -103.3609
20.8971 | -103.3513
20.9008 | -103.3422
20.9045 | -103.3333
20.9081 | -103.326
20.9118 | -103.3199
20.9155 | -103.3136
20.9191 | -103.3077
20.9228 | -103.3057
20.9265 | -103.3046
20.9301 | -103.3043
20.9338 | -103.3051
20.9375 | -103.3065
20.9411 | -103.3085
20.9448 | -103.3095
20.9485 | -103.314
20.9521 | -103.3172
20.9558 | -103.3212
20.9595 | -103.3266
20.9631 | -103.3333
20.9668 | -103.3409
20.9704 | -103.3489
20.9741 | -103.3584
20.9778 | -103.3662
20.9814 | -103.3751
20.9851 | -103.3865
20.9888 | -103.3974
20.9924 | -103.4069
20.9961 | -103.4174
20.9998 | -103.4293
21.0034 | -103.4408
21.0071 | -103.4526
21.0108 | -103.4654
21.0144 | -103.4771
21.0181 | -103.4881
21.0218 | -103.5007
21.0254 | -103.5146
21.0291 | -103.5267
21.0328 | -103.5371
21.0364 | -103.5472
21.0401 | -103.5569
21.0438 | -103.5664
21.0474 | -103.5761
21.0511 | -103.5855
21.0548 | -103.5942
21.0584 | -103.6016
21.0621 | -103.6089
21.0658 | -103.617
21.0694 | -103.6245
21.0731 | -103.6316
21.0767 | -103.6377
21.0804 | -103.6423
21.0841 | -103.6481
21.0877 | -103.6541
21.0914 | -103.6603
21.0951 | -103.6664
21.0987 | -103.673
21.1024 | -103.6798
21.1061 | -103.6861
21.1097 | -103.693
21.1134 | -103.7007
21.1171 | -103.7095
21.1207 | -103.7185
21.1244 | -103.7265
21.1281 | -103.7324
21.1317 | -103.7369
21.1354 | -103.7417
21.1391 | -103.7464
21.1427 | -103.7492
21.1464 | -103.7515
21.1501 | -103.7535
21.1537 | -103.7556
21.1574 | -103.7595
21.1611 | -103.7613
21.1647 | -103.7621
21.1684 | -103.7647
21.172 | -103.7659
21.1757 | -103.7641
21.1794 | -103.7635
21.183 | -103.7626
21.1867 | -103.7609
21.1904 | -103.7618
21.194 | -103.7622
21.1977 | -103.7622
21.2014 | -103.763
21.205 | -103.7629
21.2087 | -103.7633
21.2124 | -103.7638
21.216 | -103.766
21.2197 | -103.768
21.2234 | -103.7689
21.227 | -103.7718
21.2307 | -103.776
21.2344 | -103.7775
21.238 | -103.7795
21.2417 | -103.7836
21.2454 | -103.7902
21.249 | -103.7938
21.2527 | -103.7977
21.2564 | -103.8019
21.26 | -103.8062
21.2637 | -103.8101
21.2674 | -103.8137
21.271 | -103.8172
21.2747 | -103.82
21.2784 | -103.8237
21.282 | -103.8281
21.2857 | -103.8311
21.2893 | -103.8338
21.293 | -103.8379
21.2967 | -103.8423
21.3003 | -103.8457
21.304 | -103.8506
21.3077 | -103.8546
21.3113 | -103.8573
21.315 | -103.8592
21.3187 | -103.8612
21.3223 | -103.8618
21.326 | -103.8614
21.3297 | -103.8634
21.3333 | -103.8673
21.337 | -103.871
21.3407 | -103.8745
21.3443 | -103.8785
21.348 | -103.8822
21.3517 | -103.8846
21.3553 | -103.885
21.359 | -103.8851
21.3627 | -103.8856
21.3663 | -103.8862
21.37 | -103.8867
21.3737 | -103.8866
21.3773 | -103.8858
21.381 | -103.8845
21.3847 | -103.883
21.3883 | -103.8815
21.392 | -103.8818
21.3956 | -103.881
21.3993 | -103.8776
21.403 | -103.8758
21.4066 | -103.8732
21.4103 | -103.8698
21.414 | -103.8677
21.4176 | -103.8658
21.4213 | -103.863
21.425 | -103.8603
21.4286 | -103.8576
21.4323 | -103.8543
21.436 | -103.8509
21.4396 | -103.8487
21.4433 | -103.8479
21.447 | -103.8472
21.4506 | -103.846
21.4543 | -103.8444
21.458 | -103.8432
21.4616 | -103.8415
21.4653 | -103.8405
21.469 | -103.8392
21.4726 | -103.8367
21.4763 | -103.835
21.48 | -103.8333
21.4836 | -103.8308
21.4873 | -103.8314
21.491 | -103.8318
21.4946 | -103.8264
21.4983 | -103.8214
21.5019 | -103.8197
21.5056 | -103.8193
21.5093 | -103.8192
21.5129 | -103.8167
21.5166 | -103.8127
21.5203 | -103.8061
21.5239 | -103.7987
21.5276 | -103.7899
21.5313 | -103.7783
21.5349 | -103.7624
21.5386 | -103.7448
21.5423 | -103.7258
21.5459 | -103.7051
21.5496 | -103.682
21.5533 | -103.6589
21.5569 | -103.6347
21.5606 | -103.6064
21.5643 | -103.5763
21.5679 | -103.5453
21.5716 | -103.5113
21.5753 | -103.4776
21.5789 | -103.4422
21.5826 | -103.4072
21.5863 | -103.3687
21.5899 | -103.3275
21.5936 | -103.2856
21.5973 | -103.2428
21.6009 | -103.1988
21.6046 | -103.1534
21.6082 | -103.1063
21.6119 | -103.0588
21.6156 | -103.0109
21.6192 | -102.962
21.6229 | -102.9114
21.6266 | -102.86
21.6302 | -102.8092
21.6339 | -102.7592
21.6376 | -102.7086
21.6412 | -102.6591
21.6449 | -102.6098
21.6486 | -102.5615
21.6522 | -102.5151
21.6559 | -102.4708
21.6596 | -102.4289
21.6632 | -102.3882
21.6669 | -102.3492
21.6706 | -102.3127
21.6742 | -102.2776
21.6779 | -102.2461
21.6816 | -102.2377
21.6852 | -102.2128
21.6889 | -102.1896
21.6926 | -102.1699
21.6962 | -102.1535
21.6999 | -102.1396
21.7036 | -102.1289
21.7072 | -102.1228
21.7109 | -102.1194
21.7145 | -102.1199
21.7182 | -102.1244
21.7219 | -102.131
21.7255 | -102.1406
21.7292 | -102.153
21.7329 | -102.1688
21.7365 | -102.1873
21.7402 | -102.208
21.7439 | -102.2304
21.7475 | -102.2549
21.7512 | -102.2832
21.7549 | -102.313
21.7585 | -102.3438
21.7622 | -102.3765
21.7659 | -102.4119
21.7695 | -102.4497
21.7732 | -102.4894
21.7769 | -102.5301
21.7805 | -102.5714
21.7842 | -102.6129
21.7879 | -102.6557
21.7915 | -102.6983
21.7952 | -102.7394
21.7989 | -102.7796
21.8025 | -102.8192
21.8062 | -102.8597
21.8099 | -102.898
21.8135 | -102.932
21.8172 | -102.9646
21.8208 | -102.9985
21.8245 | -103.0321
21.8282 | -103.0638
21.8318 | -103.0929
21.8355 | -103.1217
21.8392 | -103.1485
21.8428 | -103.1746
21.8465 | -103.1994
21.8502 | -103.2224
21.8538 | -103.2419
21.8575 | -103.2575
21.8612 | -103.2682
21.8648 | -103.2749
21.8685 | -103.2799
21.8722 | -103.2806
21.8758 | -103.2773
21.8795 | -103.274
21.8832 | -103.2657
21.8868 | -103.253
21.8905 | -103.2393
21.8942 | -103.2233
21.8978 | -103.2052
21.9015 | -103.1863
21.9052 | -103.1648
21.9088 | -103.1394
21.9125 | -103.1121
21.9162 | -103.0847
21.9198 | -103.0572
21.9235 | -103.0292
21.9271 | -103.0014
21.9308 | -102.9731
21.9345 | -102.944
21.9381 | -102.9134
21.9418 | -102.8801
21.9455 | -102.8455
21.9491 | -102.8128
21.9528 | -102.7826
21.9565 | -102.7536
21.9601 | -102.727
21.9638 | -102.6999
21.9675 | -102.6727
21.9711 | -102.6478
21.9748 | -102.6242
21.9785 | -102.6005
21.9821 | -102.5781
21.9858 | -102.5568
21.9895 | -102.5374
21.9931 | -102.52
21.9968 | -102.5047
22.0005 | -102.4933
22.0041 | -102.4821
22.0078 | -102.4735
22.0115 | -102.467
22.0151 | -102.4623
22.0188 | -102.4605
22.0225 | -102.459
22.0261 | -102.4596
22.0298 | -102.4634
22.0334 | -102.4685
22.0371 | -102.4766
22.0408 | -102.4875
22.0444 | -102.5002
22.0481 | -102.5157
22.0518 | -102.5331
22.0554 | -102.5527
22.0591 | -102.574
22.0628 | -102.596
22.0664 | -102.6199
22.0701 | -102.6474
22.0738 | -102.6769
22.0774 | -102.7056
22.0811 | -102.7343
22.0848 | -102.7644
22.0884 | -102.7955
22.0921 | -102.829
22.0958 | -102.8642
22.0994 | -102.8999
22.1031 | -102.9346
22.1068 | -102.9702
22.1104 | -103.0087
22.1141 | -103.0494
22.1178 | -103.0886
22.1214 | -103.1254
22.1251 | -103.1625
22.1288 | -103.1998
22.1324 | -103.2349
22.1361 | -103.2706
22.1397 | -103.3062
22.1434 | -103.3398
22.1471 | -103.3717
22.1507 | -103.4033
22.1544 | -103.433
22.1581 | -103.46
22.1617 | -103.4845
22.1654 | -103.508
22.1691 | -103.5315
22.1727 | -103.5544
22.1764 | -103.5784
22.1801 | -103.5994
22.1837 | -103.6192
22.1874 | -103.6374
22.1911 | -103.6513
22.1947 | -103.6652
22.1984 | -103.6786
22.2021 | -103.6883
22.2057 | -103.6987
22.2094 | -103.6971
22.2131 | -103.7067
22.2167 | -103.7143
22.2204 | -103.7202
22.2241 | -103.7254
22.2277 | -103.7286
22.2314 | -103.7298
22.2351 | -103.7305
22.2387 | -103.7307
22.2424 | -103.7295
22.246 | -103.7279
22.2497 | -103.727
22.2534 | -103.7255
22.257 | -103.724
22.2607 | -103.7227
22.2644 | -103.7167
22.268 | -103.7085
22.2717 | -103.7005
22.2754 | -103.688
22.279 | -103.6756
22.2827 | -103.6651
22.2864 | -103.6551
22.29 | -103.6439
22.2937 | -103.6317
22.2974 | -103.6193
22.301 | -103.6044
22.3047 | -103.5849
22.3084 | -103.5651
22.312 | -103.5445
22.3157 | -103.521
22.3194 | -103.4964
22.323 | -103.4709
22.3267 | -103.4439
22.3304 | -103.4149
22.334 | -103.3839
22.3377 | -103.3554
22.3414 | -103.3253
22.345 | -103.2912
22.3487 | -103.2571
22.3523 | -103.2233
22.356 | -103.1878
22.3597 | -103.1507
22.3633 | -103.114
22.367 | -103.077
22.3707 | -103.0387
22.3743 | -102.9981
22.378 | -102.956
22.3817 | -102.9138
22.3853 | -102.8701
22.389 | -102.8252
22.3927 | -102.7823
22.3963 | -102.7394
22.4 | -102.6957
22.4037 | -102.6505
22.4073 | -102.6042
22.411 | -102.5587
22.4147 | -102.515
22.4183 | -102.4738
22.422 | -102.4325
22.4257 | -102.3892
22.4293 | -102.3474
22.433 | -102.3057
22.4367 | -102.267
22.4403 | -102.2304
22.444 | -102.1942
22.4477 | -102.1593
22.4513 | -102.1249
22.455 | -102.094
22.4586 | -102.0656
22.4623 | -102.0379
22.466 | -102.0112
22.4696 | -101.9862
22.4733 | -101.9646
22.477 | -101.9449
22.4806 | -101.9269
22.4843 | -101.912
22.488 | -101.8997
22.4916 | -101.8876
22.4953 | -101.8763
22.499 | -101.8666
22.5026 | -101.8589
22.5063 | -101.853
22.51 | -101.8506
22.5136 | -101.8507
22.5173 | -101.8518
22.521 | -101.8533
22.5246 | -101.8571
22.5283 | -101.8627
22.532 | -101.8684
22.5356 | -101.8745
22.5393 | -101.8817
22.543 | -101.8906
22.5466 | -101.9016
22.5503 | -101.9141
22.554 | -101.925
22.5576 | -101.9362
22.5613 | -101.9486
22.5649 | -101.9621
22.5686 | -101.9769
22.5723 | -101.9927
22.5759 | -102.0071
22.5796 | -102.0222
22.5833 | -102.0378
22.5869 | -102.0524
22.5906 | -102.0663
22.5943 | -102.0822
22.5979 | -102.0989
22.6016 | -102.1141
22.6053 | -102.1273
22.6089 | -102.1393
22.6126 | -102.1506
22.6163 | -102.1637
22.6199 | -102.1768
22.6236 | -102.1902
22.6273 | -102.2019
22.6309 | -102.2136
22.6346 | -102.2242
22.6383 | -102.2355
22.6419 | -102.2458
22.6456 | -102.2567
22.6493 | -102.266
22.6529 | -102.273
22.6566 | -102.2805
22.6603 | -102.2882
22.6639 | -102.2964
22.6676 | -102.3052
22.6712 | -102.3132
22.6749 | -102.3226
22.6786 | -102.3325
22.6822 | -102.342
22.6859 | -102.3525
22.6896 | -102.3631
22.6932 | -102.3743
22.6969 | -102.3879
22.7006 | -102.4014
22.7042 | -102.415
22.7079 | -102.4312
22.7116 | -102.4472
22.7152 | -102.4636
22.7189 | -102.4812
22.7226 | -102.5024
22.7262 | -102.5225
22.7299 | -102.5425
22.7336 | -102.5619
22.7372 | -102.5834
22.7409 | -102.6036
22.7446 | -102.6252
22.7482 | -102.6481
22.7519 | -102.6708
22.7556 | -102.6942
22.7592 | -102.7161
22.7629 | -102.7386
22.7666 | -102.7596
22.7702 | -102.7821
22.7739 | -102.8059
22.7775 | -102.8261
22.7812 | -102.8458
22.7849 | -102.8659
22.7885 | -102.8878
22.7922 | -102.9111
22.7959 | -102.9303
22.7995 | -102.9498
22.8032 | -102.9705
22.8069 | -102.9895
22.8105 | -103.0061
22.8142 | -103.0243
22.8179 | -103.0447
22.8215 | -103.0646
22.8252 | -103.0846
22.8289 | -103.1034
22.8325 | -103.1188
22.8362 | -103.1318
22.8399 | -103.1458
22.8435 | -103.1598
22.8472 | -103.1745
22.8509 | -103.189
22.8545 | -103.2028
22.8582 | -103.2154
22.8619 | -103.2267
22.8655 | -103.237
22.8692 | -103.2456
22.8729 | -103.2515
22.8765 | -103.2589
22.8802 | -103.2654
22.8838 | -103.2697
22.8875 | -103.2752
22.8912 | -103.2769
22.8948 | -103.2774
22.8985 | -103.2796
22.9022 | -103.279
22.9058 | -103.2786
22.9095 | -103.2802
22.9132 | -103.2811
22.9168 | -103.2808
22.9205 | -103.2793
22.9242 | -103.2765
22.9278 | -103.2726
22.9315 | -103.2695
22.9352 | -103.2665
22.9388 | -103.262
22.9425 | -103.2579
22.9462 | -103.2545
22.9498 | -103.2509
22.9535 | -103.2485
22.9572 | -103.2468
22.9608 | -103.2422
22.9645 | -103.2358
22.9682 | -103.2306
22.9718 | -103.2267
22.9755 | -103.2242
22.9792 | -103.2178
22.9828 | -103.2129
22.9865 | -103.2087
22.9901 | -103.2034
22.9938 | -103.1979
22.9975 | -103.1921
23.0011 | -103.1874
23.0048 | -103.1833
23.0085 | -103.1775
23.0121 | -103.172
23.0158 | -103.1669
23.0195 | -103.1606
23.0231 | -103.1533
23.0268 | -103.1479
23.0305 | -103.1418
23.0341 | -103.1368
23.0378 | -103.1312
23.0415 | -103.126
23.0451 | -103.118
23.0488 | -103.1133
23.0525 | -103.1077
23.0561 | -103.1005
23.0598 | -103.0924
23.0635 | -103.0821
23.0671 | -103.073
23.0708 | -103.0618
23.0745 | -103.0506
23.0781 | -103.0383
23.0818 | -103.0264
23.0855 | -103.0171
23.0891 | -103.0077
23.0928 | -102.9981
23.0964 | -102.9837
23.1001 | -102.9661
23.1038 | -102.951
23.1074 | -102.9381
23.1111 | -102.9257
23.1148 | -102.9099
23.1184 | -102.894
23.1221 | -102.8801
23.1258 | -102.8645
23.1294 | -102.8475
23.1331 | -102.8317
23.1368 | -102.8163
23.1404 | -102.7989
23.1441 | -102.7808
23.1478 | -102.7626
23.1514 | -102.7443
23.1551 | -102.7287
23.1588 | -102.7129
23.1624 | -102.6982
23.1661 | -102.6805
23.1698 | -102.6626
23.1734 | -102.6459
23.1771 | -102.6284
23.1808 | -102.6119
23.1844 | -102.5959
23.1881 | -102.5816
23.1918 | -102.5676
23.1954 | -102.5532
23.1991 | -102.5392
23.2027 | -102.5258
23.2064 | -102.5136
23.2101 | -102.5012
23.2137 | -102.4863
23.2174 | -102.4731
23.2211 | -102.4615
23.2247 | -102.4513
23.2284 | -102.4424
23.2321 | -102.4334
23.2357 | -102.4234
23.2394 | -102.4139
23.2431 | -102.4061
23.2467 | -102.3975
23.2504 | -102.3898
23.2541 | -102.3832
23.2577 | -102.3774
23.2614 | -102.3738
23.2651 | -102.3749
23.2687 | -102.3708
23.2724 | -102.3675
23.2761 | -102.3651
23.2797 | -102.3629
23.2834 | -102.3593
23.2871 | -102.3573
23.2907 | -102.358
23.2944 | -102.3622
23.2981 | -102.365
23.3017 | -102.3688
23.3054 | -102.3732
23.309 | -102.3793
23.3127 | -102.3851
23.3164 | -102.3915
23.32 | -102.3996
23.3237 | -102.4088
23.3274 | -102.4175
23.331 | -102.4274
23.3347 | -102.4392
23.3384 | -102.4529
23.342 | -102.4679
23.3457 | -102.4843
23.3494 | -102.5009
23.353 | -102.5189
23.3567 | -102.5358
23.3604 | -102.5529
23.364 | -102.5687
23.3677 | -102.587
23.3714 | -102.6064
23.375 | -102.6276
23.3787 | -102.6493
23.3824 | -102.6719
23.386 | -102.696
23.3897 | -102.7226
23.3934 | -102.7501
23.397 | -102.7738
23.4007 | -102.8002
23.4044 | -102.8269
23.408 | -102.8543
23.4117 | -102.8823
23.4153 | -102.909
23.419 | -102.9344
23.4227 | -102.9606
23.4263 | -102.9901
23.43 | -103.0213
23.4337 | -103.0503
23.4373 | -103.0763
23.441 | -103.1025
23.4447 | -103.1307
23.4483 | -103.1579
23.452 | -103.1847
23.4557 | -103.2107
23.4593 | -103.2353
23.463 | -103.2591
23.4667 | -103.2841
23.4703 | -103.3108
23.474 | -103.336
23.4777 | -103.3608
23.4813 | -103.384
23.485 | -103.4044
23.4887 | -103.425
23.4923 | -103.445
23.496 | -103.466
23.4997 | -103.4872
23.5033 | -103.506
23.507 | -103.5247
23.5107 | -103.543
23.5143 | -103.5613
23.518 | -103.5769
23.5216 | -103.5917
23.5253 | -103.6079
23.529 | -103.6226
23.5326 | -103.6347
23.5363 | -103.6463
23.54 | -103.6582
23.5436 | -103.672
23.5473 | -103.6852
23.551 | -103.6982
23.5546 | -103.7082
23.5583 | -103.7174
23.562 | -103.7271
23.5656 | -103.7372
23.5693 | -103.7471
23.573 | -103.7559
23.5766 | -103.7677
23.5803 | -103.7803
23.584 | -103.7914
23.5876 | -103.8023
23.5913 | -103.8115
23.595 | -103.8203
23.5986 | -103.8309
23.6023 | -103.8414
23.606 | -103.8498
23.6096 | -103.8577
23.6133 | -103.8662
23.617 | -103.875
23.6206 | -103.884
23.6243 | -103.8931
23.6279 | -103.9005
23.6316 | -103.9095
23.6353 | -103.9179
23.6389 | -103.926
23.6426 | -103.9347
23.6463 | -103.9419
23.6499 | -103.9493
23.6536 | -103.9547
23.6573 | -103.9589
23.6609 | -103.9655
23.6646 | -103.9712
23.6683 | -103.9771
23.6719 | -103.9835
23.6756 | -103.9901
23.6793 | -103.9965
23.6829 | -104.0031
23.6866 | -104.0095
23.6903 | -104.0132
23.6939 | -104.019
23.6976 | -104.0276
23.7013 | -104.0355
23.7049 | -104.0408
23.7086 | -104.0462
23.7123 | -104.0506
23.7159 | -104.0541
23.7196 | -104.0578
23.7233 | -104.0616
23.7269 | -104.0666
23.7306 | -104.0734
23.7342 | -104.0797
23.7379 | -104.0863
23.7416 | -104.0928
23.7452 | -104.1
23.7489 | -104.106
23.7526 | -104.1114
23.7562 | -104.1183
23.7599 | -104.124
23.7636 | -104.1292
23.7672 | -104.1346
23.7709 | -104.139
23.7746 | -104.1412
23.7782 | -104.1432
23.7819 | -104.1472
23.7856 | -104.1514
23.7892 | -104.1567
23.7929 | -104.1605
23.7966 | -104.161
23.8002 | -104.1618
23.8039 | -104.1633
23.8076 | -104.1657
23.8112 | -104.1669
23.8149 | -104.1646
23.8186 | -104.1632
23.8222 | -104.1636
23.8259 | -104.1634
23.8296 | -104.1633
23.8332 | -104.1638
23.8369 | -104.1628
23.8406 | -104.1615
23.8442 | -104.158
23.8479 | -104.1546
23.8515 | -104.152
23.8552 | -104.1514
23.8589 | -104.1499
23.8625 | -104.1481
23.8662 | -104.1463
23.8699 | -104.1464
23.8735 | -104.1448
23.8772 | -104.1434
23.8809 | -104.1408
23.8845 | -104.1386
23.8882 | -104.1343
23.8919 | -104.1314
23.8955 | -104.1292
23.8992 | -104.1244
23.9029 | -104.1216
23.9065 | -104.1193
23.9102 | -104.1162
23.9139 | -104.1116
23.9175 | -104.1074
23.9212 | -104.1048
23.9249 | -104.1031
23.9285 | -104.0999
23.9322 | -104.0963
23.9359 | -104.0912
23.9395 | -104.0873
23.9432 | -104.083
23.9469 | -104.0805
23.9505 | -104.0778
23.9542 | -104.076
23.9578 | -104.0734
23.9615 | -104.0696
23.9652 | -104.066
23.9688 | -104.0627
23.9725 | -104.0611
23.9762 | -104.0585
23.9798 | -104.0562
23.9835 | -104.0531
23.9872 | -104.0512
23.9908 | -104.0528
23.9945 | -104.0523
23.9982 | -104.049
24.0018 | -104.0476
24.0055 | -104.046
24.0092 | -104.044
24.0128 | -104.0419
24.0165 | -104.0391
24.0202 | -104.0365
24.0238 | -104.0343
24.0275 | -104.0352
24.0312 | -104.0341
24.0348 | -104.0316
24.0385 | -104.0271
24.0422 | -104.0259
24.0458 | -104.0253
24.0495 | -104.0244
24.0531 | -104.0234
24.0568 | -104.0226
24.0605 | -104.0205
24.0641 | -104.0167
24.0678 | -104.012
24.0715 | -104.0077
24.0751 | -104.0029
24.0788 | -103.9985
24.0825 | -103.9932
24.0861 | -103.9909
24.0898 | -103.9888
24.0935 | -103.9862
24.0971 | -103.9831
24.1008 | -103.9802
24.1045 | -103.9775
24.1081 | -103.9758
24.1118 | -103.9733
24.1155 | -103.97
24.1191 | -103.9662
24.1228 | -103.9619
24.1265 | -103.9573
24.1301 | -103.9526
24.1338 | -103.9468
24.1375 | -103.9407
24.1411 | -103.9352
24.1448 | -103.9293
24.1485 | -103.9222
24.1521 | -103.9159
24.1558 | -103.9117
24.1594 | -103.9054
24.1631 | -103.8981
24.1668 | -103.8909
24.1704 | -103.884
24.1741 | -103.8771
24.1778 | -103.8705
24.1814 | -103.8624
24.1851 | -103.8564
24.1888 | -103.8496
24.1924 | -103.8429
24.1961 | -103.835
24.1998 | -103.8251
24.2034 | -103.8155
24.2071 | -103.8071
24.2108 | -103.7991
24.2144 | -103.7933
24.2181 | -103.7866
24.2218 | -103.7799
24.2254 | -103.7746
24.2291 | -103.7721
24.2328 | -103.769
24.2364 | -103.7651
24.2401 | -103.7624
24.2438 | -103.7587
24.2474 | -103.7546
24.2511 | -103.7514
24.2548 | -103.75
24.2584 | -103.7464
24.2621 | -103.7421
24.2658 | -103.7346
24.2694 | -103.7248
24.2731 | -103.715
24.2767 | -103.7044
24.2804 | -103.6887
24.2841 | -103.6667
24.2877 | -103.6482
24.2914 | -103.6354
24.2951 | -103.6214
24.2987 | -103.6083
24.3024 | -103.5944
24.3061 | -103.5786
24.3097 | -103.564
24.3134 | -103.5472
24.3171 | -103.5303
24.3207 | -103.5125
24.3244 | -103.4828
24.3281 | -103.4625
24.3317 | -103.4403
24.3354 | -103.4152
24.3391 | -103.3922
24.3427 | -103.3697
24.3464 | -103.3446
24.3501 | -103.3186
24.3537 | -103.2925
24.3574 | -103.2656
24.3611 | -103.2385
24.3647 | -103.2095
24.3684 | -103.1809
24.3721 | -103.1519
24.3757 | -103.1233
24.3794 | -103.0939
24.383 | -103.0664
24.3867 | -103.0374
24.3904 | -103.0087
24.394 | -102.9798
24.3977 | -102.9535
24.4014 | -102.929
24.405 | -102.9056
24.4087 | -102.8822
24.4124 | -102.8591
24.416 | -102.8379
24.4197 | -102.819
24.4234 | -102.8024
24.427 | -102.7879
24.4307 | -102.7738
24.4344 | -102.7614
24.438 | -102.7516
24.4417 | -102.7434
24.4454 | -102.7366
24.449 | -102.7302
24.4527 | -102.7257
24.4564 | -102.7261
24.46 | -102.728
24.4637 | -102.7313
24.4674 | -102.7362
24.471 | -102.743
24.4747 | -102.7507
24.4784 | -102.7609
24.482 | -102.772
24.4857 | -102.7846
24.4893 | -102.7977
24.493 | -102.8111
24.4967 | -102.8264
24.5003 | -102.8413
24.504 | -102.8569
24.5077 | -102.8717
24.5113 | -102.8855
24.515 | -102.9005
24.5187 | -102.9158
24.5223 | -102.9293
24.526 | -102.9442
24.5297 | -102.9579
24.5333 | -102.9704
24.537 | -102.9829
24.5407 | -102.9941
24.5443 | -103.0059
24.548 | -103.0171
24.5517 | -103.024
24.5553 | -103.0247
24.559 | -103.0206
24.5627 | -103.0111
24.5663 | -102.9991
24.57 | -102.981
24.5737 | -102.9582
24.5773 | -102.9255
24.581 | -102.8745
24.5847 | -102.7801
24.5883 | -102.6632
24.592 | -102.5308
24.5956 | -102.3842
24.5993 | -102.2324
24.603 | -102.0967
24.6066 | -101.976
24.6103 | -101.8589
24.614 | -101.7414
24.6176 | -101.6182
24.6213 | -101.4899
24.625 | -101.3577
24.6286 | -101.2272
24.6323 | -101.0879
24.636 | -100.9435
24.6396 | -100.7967
24.6433 | -100.6478
24.647 | -100.4975
24.6506 | -100.3467
24.6543 | -100.1944
24.658 | -100.0422
24.6616 | -99.8904
24.6653 | -99.7401
24.669 | -99.5913
24.6726 | -99.4445
24.6763 | -99.2999
24.68 | -99.1572
24.6836 | -99.017
24.6873 | -98.8809
24.691 | -98.749
24.6946 | -98.6219
24.6983 | -98.4992
24.7019 | -98.3814
24.7056 | -98.2689
24.7093 | -98.1622
24.7129 | -98.0617
24.7166 | -97.9671
24.7203 | -97.8788
24.7239 | -97.7968
24.7276 | -97.7215
24.7313 | -97.6529
24.7349 | -97.5908
24.7386 | -97.5357
24.7423 | -97.4882
24.7459 | -97.448
24.7496 | -97.4149
24.7533 | -97.3891
24.7569 | -97.3712
24.7606 | -97.361
24.7643 | -97.3577
24.7679 | -97.362
24.7716 | -97.3737
24.7753 | -97.3933
24.7789 | -97.4205
24.7826 | -97.4552
24.7863 | -97.4973
24.7899 | -97.5469
24.7936 | -97.604
24.7973 | -97.6689
24.8009 | -97.7413
24.8046 | -97.8199
24.8082 | -97.9052
24.8119 | -97.9968
24.8156 | -98.0953
24.8192 | -98.1999
24.8229 | -98.3115
24.8266 | -98.4284
24.8302 | -98.5497
24.8339 | -98.6763
24.8376 | -98.8075
24.8412 | -98.943
24.8449 | -99.0828
24.8486 | -99.2271
24.8522 | -99.357
24.8559 | -99.5016
24.8596 | -99.6487
24.8632 | -99.7985
24.8669 | -99.9488
24.8706 | -100.0998
24.8742 | -100.252
24.8779 | -100.4039
24.8816 | -100.5551
24.8852 | -100.7051
24.8889 | -100.8532
24.8926 | -101.0001
24.8962 | -101.1443
24.8999 | -101.284
24.9036 | -101.4223
24.9072 | -101.5588
24.9109 | -101.6911
24.9145 | -101.8207
24.9182 | -101.9457
24.9219 | -102.0665
24.9255 | -102.1823
24.9292 | -102.2972
24.9329 | -102.4176
24.9365 | -102.5617
24.9402 | -102.7069
24.9439 | -102.848
24.9475 | -102.9832
24.9512 | -103.1019
24.9549 | -103.1846
24.9585 | -103.2345
24.9622 | -103.2667
24.9659 | -103.2905
24.9695 | -103.3117
24.9732 | -103.3306
24.9769 | -103.3466
24.9805 | -103.3548
24.9842 | -103.3626
24.9879 | -103.3671
24.9915 | -103.3702
24.9952 | -103.3708
24.9989 | -103.3721
25.0025 | -103.3755
25.0062 | -103.3797
25.0099 | -103.3818
25.0135 | -103.382
25.0172 | -103.3827
25.0208 | -103.3835
25.0245 | -103.3838
25.0282 | -103.3855
25.0318 | -103.3885
25.0355 | -103.3915
25.0392 | -103.3923
25.0428 | -103.392
25.0465 | -103.3913
25.0502 | -103.3895
25.0538 | -103.3876
25.0575 | -103.3866
25.0612 | -103.3833
25.0648 | -103.3778
25.0685 | -103.3723
25.0722 | -103.3669
25.0758 | -103.3619
25.0795 | -103.3563
25.0832 | -103.3499
25.0868 | -103.3444
25.0905 | -103.3402
25.0942 | -103.335
25.0978 | -103.3287
25.1015 | -103.3215
25.1052 | -103.3157
25.1088 | -103.309
25.1125 | -103.3019
25.1162 | -103.2936
25.1198 | -103.2857
25.1235 | -103.2784
25.1271 | -103.2697
25.1308 | -103.2612
25.1345 | -103.2532
25.1381 | -103.2447
25.1418 | -103.236
25.1455 | -103.228
25.1491 | -103.2176
25.1528 | -103.207
25.1565 | -103.1978
25.1601 | -103.1892
25.1638 | -103.1799
25.1675 | -103.1701
25.1711 | -103.1608
25.1748 | -103.152
25.1785 | -103.1448
25.1821 | -103.1389
25.1858 | -103.1302
25.1895 | -103.1221
25.1931 | -103.116
25.1968 | -103.1111
25.2005 | -103.1038
25.2041 | -103.0968
25.2078 | -103.0918
25.2115 | -103.0875
25.2151 | -103.0829
25.2188 | -103.0781
25.2225 | -103.0729
25.2261 | -103.0693
25.2298 | -103.0666
25.2334 | -103.0631
25.2371 | -103.0625
25.2408 | -103.064
25.2444 | -103.0651
25.2481 | -103.0662
25.2518 | -103.0678
25.2554 | -103.0703
25.2591 | -103.0718
25.2628 | -103.0745
25.2664 | -103.0786
25.2701 | -103.0825
25.2738 | -103.0846
25.2774 | -103.0866
25.2811 | -103.0889
25.2848 | -103.0932
25.2884 | -103.0979
25.2921 | -103.102
25.2958 | -103.1062
25.2994 | -103.1095
25.3031 | -103.1127
25.3068 | -103.1172
25.3104 | -103.1188
25.3141 | -103.1205
25.3178 | -103.1246
25.3214 | -103.1281
25.3251 | -103.1309
25.3288 | -103.1343
25.3324 | -103.1362
25.3361 | -103.1386
25.3397 | -103.1394
25.3434 | -103.1415
25.3471 | -103.144
25.3507 | -103.1466
25.3544 | -103.1476
25.3581 | -103.1459
25.3617 | -103.1431
25.3654 | -103.1434
25.3691 | -103.1443
25.3727 | -103.1427
25.3764 | -103.1402
25.3801 | -103.1344
25.3837 | -103.1317
25.3874 | -103.1279
25.3911 | -103.1246
25.3947 | -103.1219
25.3984 | -103.1171
25.4021 | -103.1124
25.4057 | -103.1071
25.4094 | -103.1029
25.4131 | -103.0985
25.4167 | -103.0946
25.4204 | -103.0918
25.4241 | -103.09
25.4277 | -103.0887
25.4314 | -103.0868
25.4351 | -103.0836
25.4387 | -103.0798
25.4424 | -103.0761
25.446 | -103.073
25.4497 | -103.0696
25.4534 | -103.0669
25.457 | -103.0646
25.4607 | -103.0649
25.4644 | -103.0647
25.468 | -103.0656
25.4717 | -103.0665
25.4754 | -103.0706
25.479 | -103.0725
25.4827 | -103.0735
25.4864 | -103.0767
25.49 | -103.0795
25.4937 | -103.0836
25.4974 | -103.0877
25.501 | -103.0927
25.5047 | -103.0948
25.5084 | -103.0973
25.512 | -103.1001
25.5157 | -103.1031
25.5194 | -103.1074
25.523 | -103.1125
25.5267 | -103.1178
25.5304 | -103.1228
25.534 | -103.1266
25.5377 | -103.1306
25.5414 | -103.1355
25.545 | -103.1409
25.5487 | -103.1457
25.5523 | -103.1485
25.556 | -103.1494
25.5597 | -103.1521
25.5633 | -103.1537
25.567 | -103.1536
25.5707 | -103.1557
25.5743 | -103.1575
25.578 | -103.1575
25.5817 | -103.1587
25.5853 | -103.1581
25.589 | -103.1586
25.5927 | -103.1585
25.5963 | -103.1563
25.6 | -103.1542
25.6037 | -103.1516
25.6073 | -103.149
25.611 | -103.1471
25.6147 | -103.145
25.6183 | -103.1437
25.622 | -103.1412
25.6257 | -103.137
25.6293 | -103.1333
25.633 | -103.1325
25.6367 | -103.1316
25.6403 | -103.1296
25.644 | -103.1252
25.6477 | -103.1211
25.6513 | -103.1184
25.655 | -103.1155
25.6586 | -103.1128
25.6623 | -103.11
25.666 | -103.1067
25.6696 | -103.1064
25.6733 | -103.1046
25.677 | -103.1028
25.6806 | -103.1019
25.6843 | -103.1005
25.688 | -103.0987
25.6916 | -103.0999
25.6953 | -103.1019
25.699 | -103.1042
25.7026 | -103.104
25.7063 | -103.1035
25.71 | -103.1039
25.7136 | -103.1066
25.7173 | -103.1102
25.721 | -103.1138
25.7246 | -103.1161
25.7283 | -103.1188
25.732 | -103.1216
25.7356 | -103.1263
25.7393 | -103.1309
25.743 | -103.1343
25.7466 | -103.1382
25.7503 | -103.1404
25.754 | -103.1415
25.7576 | -103.1456
25.7613 | -103.1522
25.7649 | -103.1587
25.7686 | -103.1652
25.7723 | -103.171
25.7759 | -103.1774
25.7796 | -103.1836
25.7833 | -103.1902
25.7869 | -103.1953
25.7906 | -103.2001
25.7943 | -103.2053
25.7979 | -103.2111
25.8016 | -103.2186
25.8053 | -103.2267
25.8089 | -103.2336
25.8126 | -103.2381
25.8163 | -103.2427
25.8199 | -103.2473
25.8236 | -103.2503
25.8273 | -103.253
25.8309 | -103.2558
25.8346 | -103.2585
25.8383 | -103.2593
25.8419 | -103.2619
25.8456 | -103.264
25.8493 | -103.2643
25.8529 | -103.2668
25.8566 | -103.2693
25.8603 | -103.2708
25.8639 | -103.2711
25.8676 | -103.2729
25.8712 | -103.2756
25.8749 | -103.2768
25.8786 | -103.2766
25.8822 | -103.2758
25.8859 | -103.275
25.8896 | -103.2755
25.8932 | -103.276
25.8969 | -103.2741
25.9006 | -103.2706
25.9042 | -103.269
25.9079 | -103.2688
25.9116 | -103.2684
25.9152 | -103.2677
25.9189 | -103.2659
25.9226 | -103.2647
25.9262 | -103.2623
25.9299 | -103.2598
25.9336 | -103.2597
25.9372 | -103.2592
25.9409 | -103.2599
25.9446 | -103.2614
25.9482 | -103.2621
25.9519 | -103.2624
25.9556 | -103.2623
25.9592 | -103.2633
25.9629 | -103.2642
25.9666 | -103.2647
25.9702 | -103.2653
25.9739 | -103.2659
25.9775 | -103.2676
25.9812 | -103.2704
25.9849 | -103.2749
25.9885 | -103.2798
25.9922 | -103.286
25.9959 | -103.2923
25.9995 | -103.2961
26.0032 | -103.3019
26.0069 | -103.3069
26.0105 | -103.3128
26.0142 | -103.3196
26.0179 | -103.3261
26.0215 | -103.333
26.0252 | -103.3383
26.0289 | -103.3432
26.0325 | -103.3492
26.0362 | -103.3547
26.0399 | -103.3613
26.0435 | -103.3671
26.0472 | -103.3724
26.0509 | -103.3805
26.0545 | -103.3885
26.0582 | -103.3947
26.0619 | -103.4012
26.0655 | -103.406
26.0692 | -103.411
26.0729 | -103.4186
26.0765 | -103.4247
26.0802 | -103.4296
26.0838 | -103.4353
26.0875 | -103.4413
26.0912 | -103.4481
26.0948 | -103.4533
26.0985 | -103.4592
26.1022 | -103.4629
26.1058 | -103.4658
26.1095 | -103.4714
26.1132 | -103.4756
26.1168 | -103.4813
26.1205 | -103.4873
26.1242 | -103.4887
26.1278 | -103.4884
26.1315 | -103.4883
26.1352 | -103.488
26.1388 | -103.4879
26.1425 | -103.4874
26.1462 | -103.4856
26.1498 | -103.4843
26.1535 | -103.4847
26.1572 | -103.4838
26.1608 | -103.4828
26.1645 | -103.4821
26.1682 | -103.4808
26.1718 | -103.4791
26.1755 | -103.4774
26.1792 | -103.4774
26.1828 | -103.476
26.1865 | -103.4747
26.1901 | -103.4734
26.1938 | -103.4736
26.1975 | -103.4737
26.2011 | -103.4722
26.2048 | -103.4706
26.2085 | -103.4694
26.2121 | -103.4686
26.2158 | -103.4694
26.2195 | -103.4682
26.2231 | -103.4649
26.2268 | -103.4634
26.2305 | -103.4642
26.2341 | -103.4646
26.2378 | -103.4642
26.2415 | -103.4636
26.2451 | -103.4653
26.2488 | -103.4665
26.2525 | -103.4679
26.2561 | -103.4677
26.2598 | -103.468
26.2635 | -103.4715
26.2671 | -103.4755
26.2708 | -103.4779
26.2745 | -103.48
26.2781 | -103.4836
26.2818 | -103.4865
26.2855 | -103.489
26.2891 | -103.4931
26.2928 | -103.4969
26.2964 | -103.5015
26.3001 | -103.506
26.3038 | -103.5092
26.3074 | -103.5117
26.3111 | -103.5137
26.3148 | -103.5154
26.3184 | -103.5179
26.3221 | -103.5204
26.3258 | -103.5225
26.3294 | -103.5241
26.3331 | -103.5252
26.3368 | -103.5253
26.3404 | -103.5277
26.3441 | -103.5301
26.3478 | -103.5306
26.3514 | -103.5314
26.3551 | -103.5309
26.3588 | -103.5308
26.3624 | -103.53
26.3661 | -103.5296
26.3698 | -103.5289
26.3734 | -103.527
26.3771 | -103.5267
26.3808 | -103.528
26.3844 | -103.5264
26.3881 | -103.5244
26.3918 | -103.5237
26.3954 | -103.5241
26.3991 | -103.5243
26.4027 | -103.5214
26.4064 | -103.5181
26.4101 | -103.5138
26.4137 | -103.5112
26.4174 | -103.5092
26.4211 | -103.5045
26.4247 | -103.5007
26.4284 | -103.4983
26.4321 | -103.4947
26.4357 | -103.4919
26.4394 | -103.4856
26.4431 | -103.4837
26.4467 | -103.4802
26.4504 | -103.4764
26.4541 | -103.4709
26.4577 | -103.4649
26.4614 | -103.4597
26.4651 | -103.4533
26.4687 | -103.45
26.4724 | -103.4461
26.4761 | -103.4422
26.4797 | -103.438
26.4834 | -103.4332
26.4871 | -103.4283
26.4907 | -103.4218
26.4944 | -103.4163
26.4981 | -103.4122
26.5017 | -103.4086
26.5054 | -103.4046
26.509 | -103.401
26.5127 | -103.3977
26.5164 | -103.3929
26.52 | -103.3903
26.5237 | -103.387
26.5274 | -103.3842
26.531 | -103.3813
26.5347 | -103.3764
26.5384 | -103.3705
26.542 | -103.3676
26.5457 | -103.3662
26.5494 | -103.3652
26.553 | -103.3633
26.5567 | -103.3623
26.5604 | -103.3623
26.564 | -103.3626
26.5677 | -103.3615
26.5714 | -103.3614
26.575 | -103.3629
26.5787 | -103.3649
26.5824 | -103.3663
26.586 | -103.3674
26.5897 | -103.3679
26.5934 | -103.3702
26.597 | -103.3737
26.6007 | -103.3752
26.6044 | -103.3745
26.608 | -103.3767
26.6117 | -103.3798
26.6153 | -103.3825
26.619 | -103.383
26.6227 | -103.3854
26.6263 | -103.386
26.63 | -103.3889
26.6337 | -103.3925
26.6373 | -103.3954
26.641 | -103.3983
26.6447 | -103.4015
26.6483 | -103.4062
26.652 | -103.409
26.6557 | -103.411
26.6593 | -103.4139
26.663 | -103.4152
26.6667 | -103.4158
26.6703 | -103.418
26.674 | -103.4219
26.6777 | -103.4248
26.6813 | -103.4267
26.685 | -103.4289
26.6887 | -103.4329
26.6923 | -103.4344
26.696 | -103.4357
26.6997 | -103.437
26.7033 | -103.4384
26.707 | -103.4394
26.7107 | -103.441
26.7143 | -103.4411
26.718 | -103.4404
26.7216 | -103.4408
26.7253 | -103.4395
26.729 | -103.437
26.7326 | -103.4339
26.7363 | -103.4325
26.74 | -103.4304
26.7436 | -103.428
26.7473 | -103.4253
26.751 | -103.4226
26.7546 | -103.4185
26.7583 | -103.4155
26.762 | -103.4121
26.7656 | -103.4078
26.7693 | -103.4023
26.773 | -103.398
26.7766 | -103.394
26.7803 | -103.3881
26.784 | -103.3827
26.7876 | -103.3772
26.7913 | -103.3716
26.795 | -103.3658
26.7986 | -103.3573
26.8023 | -103.3483
26.806 | -103.34
26.8096 | -103.3322
26.8133 | -103.3238
26.817 | -103.3173
26.8206 | -103.3094
26.8243 | -103.3003
26.828 | -103.2935
26.8316 | -103.2879
26.8353 | -103.2816
26.8389 | -103.2742
26.8426 | -103.2679
26.8463 | -103.2597
26.8499 | -103.2511
26.8536 | -103.2434
26.8573 | -103.2371
26.8609 | -103.2303
26.8646 | -103.224
26.8683 | -103.2187
26.8719 | -103.2128
26.8756 | -103.2071
26.8793 | -103.2016
26.8829 | -103.1975
26.8866 | -103.1911
26.8903 | -103.1848
26.8939 | -103.1818
26.8976 | -103.1773
26.9013 | -103.173
26.9049 | -103.1684
26.9086 | -103.1632
26.9123 | -103.1611
26.9159 | -103.1597
26.9196 | -103.1562
26.9233 | -103.1539
26.9269 | -103.1522
26.9306 | -103.1512
26.9342 | -103.1507
26.9379 | -103.1488
26.9416 | -103.1472
26.9452 | -103.1455
26.9489 | -103.146
26.9526 | -103.1471
26.9562 | -103.1475
26.9599 | -103.1476
26.9636 | -103.1492
26.9672 | -103.1523
26.9709 | -103.1535
26.9746 | -103.1523
26.9782 | -103.1512
26.9819 | -103.1519
26.9856 | -103.1521
26.9892 | -103.1511
26.9929 | -103.1501
26.9966 | -103.1496
27.0002 | -103.1488
27.0039 | -103.1485
27.0076 | -103.1496
27.0112 | -103.1475
27.0149 | -103.145
27.0186 | -103.1406
27.0222 | -103.1368
27.0259 | -103.1333
27.0296 | -103.1294
27.0332 | -103.1264
27.0369 | -103.1236
27.0405 | -103.1218
27.0442 | -103.1207
27.0479 | -103.1196
27.0515 | -103.1164
27.0552 | -103.112
27.0589 | -103.1078
27.0625 | -103.1037
27.0662 | -103.0995
27.0699 | -103.0953
27.0735 | -103.0927
27.0772 | -103.0901
27.0809 | -103.0857
27.0845 | -103.0812
27.0882 | -103.0783
27.0919 | -103.0744
27.0955 | -103.0705
27.0992 | -103.0658
27.1029 | -103.0608
27.1065 | -103.0576
27.1102 | -103.0539
27.1139 | -103.0496
27.1175 | -103.0442
27.1212 | -103.0387
27.1249 | -103.0339
27.1285 | -103.0293
27.1322 | -103.0253
27.1359 | -103.0211
27.1395 | -103.0178
27.1432 | -103.0161
27.1469 | -103.0132
27.1505 | -103.0123
27.1542 | -103.0096
27.1578 | -103.0082
27.1615 | -103.0076
27.1652 | -103.0077
27.1688 | -103.0073
27.1725 | -103.0046
27.1762 | -103.0032
27.1798 | -103.0048
27.1835 | -103.0069
27.1872 | -103.008
27.1908 | -103.0091
27.1945 | -103.0104
27.1982 | -103.0106
27.2018 | -103.0122
27.2055 | -103.0146
27.2092 | -103.0156
27.2128 | -103.0175
27.2165 | -103.0193
27.2202 | -103.0217
27.2238 | -103.0245
27.2275 | -103.0279
27.2312 | -103.0319
27.2348 | -103.0361
27.2385 | -103.0404
27.2422 | -103.0451
27.2458 | -103.0499
27.2495 | -103.0541
27.2532 | -103.0585
27.2568 | -103.0626
27.2605 | -103.0665
27.2641 | -103.0694
27.2678 | -103.075
27.2715 | -103.0819
27.2751 | -103.0862
27.2788 | -103.0903
27.2825 | -103.094
27.2861 | -103.0978
27.2898 | -103.1038
27.2935 | -103.1066
27.2971 | -103.111
27.3008 | -103.1149
27.3045 | -103.1167
27.3081 | -103.1196
27.3118 | -103.1216
27.3155 | -103.1225
27.3191 | -103.1237
27.3228 | -103.1243
27.3265 | -103.126
27.3301 | -103.1248
27.3338 | -103.1237
27.3375 | -103.1226
27.3411 | -103.1224
27.3448 | -103.1217
27.3485 | -103.1206
27.3521 | -103.119
27.3558 | -103.1179
27.3594 | -103.1185
27.3631 | -103.1188
27.3668 | -103.1184
27.3704 | -103.119
27.3741 | -103.1182
27.3778 | -103.1172
27.3814 | -103.1148
27.3851 | -103.1117
27.3888 | -103.1083
27.3924 | -103.1056
27.3961 | -103.1026
27.3998 | -103.0991
27.4034 | -103.0982
27.4071 | -103.097
27.4108 | -103.0967
27.4144 | -103.0956
27.4181 | -103.0926
27.4218 | -103.0897
27.4254 | -103.0868
27.4291 | -103.0851
27.4328 | -103.0815
27.4364 | -103.0801
27.4401 | -103.0802
27.4438 | -103.0788
27.4474 | -103.0796
27.4511 | -103.0803
27.4548 | -103.0801
27.4584 | -103.0798
27.4621 | -103.079
27.4657 | -103.0801
27.4694 | -103.0808
27.4731 | -103.0812
27.4767 | -103.0828
27.4804 | -103.0837
27.4841 | -103.084
27.4877 | -103.0846
27.4914 | -103.0849
27.4951 | -103.0874
27.4987 | -103.0893
27.5024 | -103.0919
27.5061 | -103.0955
27.5097 | -103.0988
27.5134 | -103.1014
27.5171 | -103.1044
27.5207 | -103.1059
27.5244 | -103.1069
27.5281 | -103.108
27.5317 | -103.1095
27.5354 | -103.1125
27.5391 | -103.1144
27.5427 | -103.1156
27.5464 | -103.1165
27.5501 | -103.1189
27.5537 | -103.1196
27.5574 | -103.1213
27.5611 | -103.1248
27.5647 | -103.1281
27.5684 | -103.1305
27.5721 | -103.132
27.5757 | -103.1334
27.5794 | -103.1345
27.583 | -103.1371
27.5867 | -103.1396
27.5904 | -103.1405
27.594 | -103.1404
27.5977 | -103.1414
27.6014 | -103.1432
27.605 | -103.145
27.6087 | -103.1453
27.6124 | -103.1437
27.616 | -103.1432
27.6197 | -103.1421
27.6234 | -103.1417
27.627 | -103.1399
27.6307 | -103.1371
27.6344 | -103.1345
27.638 | -103.1334
27.6417 | -103.1316
27.6454 | -103.131
27.649 | -103.1295
27.6527 | -103.1286
27.6564 | -103.1269
27.66 | -103.1232
27.6637 | -103.1197
27.6674 | -103.1182
27.671 | -103.1167
27.6747 | -103.1148
27.6784 | -103.1108
27.682 | -103.1068
27.6857 | -103.1014
27.6893 | -103.0967
27.693 | -103.0922
27.6967 | -103.0868
27.7003 | -103.0827
27.704 | -103.0775
27.7077 | -103.0727
27.7113 | -103.068
27.715 | -103.0631
27.7187 | -103.0594
27.7223 | -103.0548
27.726 | -103.052
27.7297 | -103.0482
27.7333 | -103.0434
27.737 | -103.039
27.7407 | -103.0346
27.7443 | -103.0302
27.748 | -103.0264
27.7517 | -103.0225
27.7553 | -103.0186
27.759 | -103.0143
27.7627 | -103.0094
27.7663 | -103.0035
27.77 | -102.9999
27.7737 | -102.9972
27.7773 | -102.9925
27.781 | -102.9879
27.7847 | -102.986
27.7883 | -102.9843
27.792 | -102.9803
27.7956 | -102.9757
27.7993 | -102.9689
27.803 | -102.9635
27.8066 | -102.961
27.8103 | -102.9585
27.814 | -102.9547
27.8176 | -102.9497
27.8213 | -102.9471
27.825 | -102.9445
27.8286 | -102.9407
27.8323 | -102.9378
27.836 | -102.9347
27.8396 | -102.933
27.8433 | -102.93
27.847 | -102.9261
27.8506 | -102.9225
27.8543 | -102.9204
27.858 | -102.9191
27.8616 | -102.9182
27.8653 | -102.9161
27.869 | -102.9145
27.8726 | -102.9146
27.8763 | -102.9145
27.88 | -102.9149
27.8836 | -102.9154
27.8873 | -102.9148
27.891 | -102.9137
27.8946 | -102.9127
27.8983 | -102.9117
27.9019 | -102.9104
27.9056 | -102.9097
27.9093 | -102.9093
27.9129 | -102.9108
27.9166 | -102.912
27.9203 | -102.9126
27.9239 | -102.9136
27.9276 | -102.9141
27.9313 | -102.9128
27.9349 | -102.9138
27.9386 | -102.9159
27.9423 | -102.9167
27.9459 | -102.9183
27.9496 | -102.9196
27.9533 | -102.9218
27.9569 | -102.9222
27.9606 | -102.9221
27.9643 | -102.922
27.9679 | -102.9235
27.9716 | -102.9258
27.9753 | -102.9268
27.9789 | -102.9264
27.9826 | -102.9264
27.9863 | -102.9284
27.9899 | -102.9305
27.9936 | -102.9323
27.9973 | -102.9329
28.0009 | -102.9358