## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 220 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 219
Hood | 02-92B/P | 157
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1012.6 | hPa |  | PRC PRO 01
Humidity | 37.9 | %/Hr
Internal temperature of calibration bench | 23.88 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1016 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.4 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.99 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3114 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 03/05/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 14/05/2018
Engineer | Anthony HUE | Approval |  | 18/05/2018
CEO | Robert GLEMOT | Approval |  | 18/05/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4768 | 57.4549
0.012209 | -35.9957 | 38.3914
0.018313 | -34.9416 | 28.4695
0.024417 | -34.4916 | 22.1035
0.030521 | -34.3906 | 18.3357
0.036626 | -34.2499 | 15.2379
0.04273 | -34.1667 | 12.9182
0.048834 | -34.1132 | 11.3517
0.054938 | -34.0395 | 10.1567
0.061043 | -33.9635 | 9.0075
0.067147 | -33.9543 | 8.1292
0.073251 | -33.9568 | 7.3701
0.079355 | -33.9767 | 6.8109
0.08546 | -33.996 | 6.3535
0.091564 | -33.9926 | 5.9826
0.097668 | -33.9575 | 5.7354
0.10377 | -33.9083 | 5.3895
0.10988 | -33.8763 | 5.0252
0.11598 | -33.867 | 4.6601
0.12209 | -33.8844 | 4.3742
0.12819 | -33.9182 | 4.1343
0.13429 | -33.9508 | 3.8888
0.1404 | -33.9682 | 3.6146
0.1465 | -33.9841 | 3.3555
0.15261 | -33.9914 | 3.1164
0.15871 | -33.9879 | 2.9215
0.16482 | -33.9735 | 2.7655
0.17092 | -33.9489 | 2.7064
0.17702 | -33.9299 | 2.6102
0.18313 | -33.9124 | 2.5281
0.18923 | -33.8997 | 2.4209
0.19534 | -33.8924 | 2.3159
0.20144 | -33.891 | 2.2005
0.20754 | -33.8935 | 2.1258
0.21365 | -33.8957 | 2.0506
0.21975 | -33.9001 | 1.9863
0.22586 | -33.9083 | 1.922
0.23196 | -33.9171 | 1.8508
0.23807 | -33.9262 | 1.7749
0.24417 | -33.9286 | 1.6824
0.25027 | -33.9323 | 1.5989
0.25638 | -33.9323 | 1.5072
0.26248 | -33.9304 | 1.4201
0.26859 | -33.928 | 1.3367
0.27469 | -33.9272 | 1.2494
0.2808 | -33.933 | 1.1846
0.2869 | -33.9375 | 1.1157
0.293 | -33.9431 | 1.058
0.29911 | -33.9495 | 1.0195
0.30521 | -33.9556 | 0.98266
0.31132 | -33.9598 | 0.96922
0.31742 | -33.9643 | 0.94724
0.32353 | -33.9678 | 0.92942
0.32963 | -33.9708 | 0.91487
0.33573 | -33.9728 | 0.90271
0.34184 | -33.9754 | 0.88708
0.34794 | -33.9787 | 0.85258
0.35405 | -33.9821 | 0.82316
0.36015 | -33.9857 | 0.79017
0.36626 | -33.9882 | 0.75361
0.37236 | -33.9907 | 0.70389
0.37846 | -33.9929 | 0.64636
0.38457 | -33.9918 | 0.58659
0.39067 | -33.9924 | 0.53606
0.39678 | -33.9906 | 0.47918
0.40288 | -33.9882 | 0.42529
0.40899 | -33.9843 | 0.37255
0.41509 | -33.9797 | 0.31954
0.42119 | -33.9746 | 0.27114
0.4273 | -33.9708 | 0.2212
0.4334 | -33.9673 | 0.17027
0.43951 | -33.9642 | 0.12008
0.44561 | -33.9627 | 0.073172
0.45172 | -33.9644 | 0.034394
0.45782 | -33.9642 | -0.0087803
0.46392 | -33.9655 | -0.048459
0.47003 | -33.9673 | -0.084444
0.47613 | -33.9694 | -0.1174
0.48224 | -33.9717 | -0.14871
0.48834 | -33.9741 | -0.17058
0.49445 | -33.9762 | -0.19599
0.50055 | -33.9788 | -0.21881
0.50665 | -33.9806 | -0.2374
0.51276 | -33.9817 | -0.25746
0.51886 | -33.9829 | -0.27483
0.52497 | -33.9818 | -0.28079
0.53107 | -33.9822 | -0.29076
0.53717 | -33.9826 | -0.30165
0.54328 | -33.9823 | -0.30622
0.54938 | -33.9823 | -0.31383
0.55549 | -33.9818 | -0.32223
0.56159 | -33.9817 | -0.33262
0.5677 | -33.9813 | -0.34238
0.5738 | -33.9809 | -0.35509
0.5799 | -33.981 | -0.36664
0.58601 | -33.9814 | -0.38246
0.59211 | -33.9826 | -0.4045
0.59822 | -33.9833 | -0.42062
0.60432 | -33.9844 | -0.44183
0.61043 | -33.9854 | -0.46262
0.61653 | -33.9867 | -0.48382
0.62263 | -33.9878 | -0.50656
0.62874 | -33.9893 | -0.52658
0.63484 | -33.9904 | -0.54734
0.64095 | -33.9917 | -0.56807
0.64705 | -33.9929 | -0.58771
0.65316 | -33.994 | -0.60708
0.65926 | -33.995 | -0.62955
0.66536 | -33.9945 | -0.64786
0.67147 | -33.9947 | -0.66401
0.67757 | -33.9953 | -0.68181
0.68368 | -33.9951 | -0.70182
0.68978 | -33.9952 | -0.71976
0.69589 | -33.9948 | -0.7381
0.70199 | -33.994 | -0.76017
0.70809 | -33.9934 | -0.78039
0.7142 | -33.9929 | -0.79972
0.7203 | -33.9921 | -0.81931
0.72641 | -33.9912 | -0.84033
0.73251 | -33.991 | -0.86121
0.73862 | -33.991 | -0.88572
0.74472 | -33.9913 | -0.90512
0.75082 | -33.9909 | -0.92631
0.75693 | -33.991 | -0.95003
0.76303 | -33.991 | -0.97268
0.76914 | -33.9916 | -1
0.77524 | -33.9919 | -1.0252
0.78135 | -33.9924 | -1.0487
0.78745 | -33.9931 | -1.0738
0.79355 | -33.9936 | -1.0986
0.79966 | -33.9947 | -1.1247
0.80576 | -33.9959 | -1.1554
0.81187 | -33.9963 | -1.1853
0.81797 | -33.997 | -1.2164
0.82408 | -33.9977 | -1.244
0.83018 | -33.9983 | -1.2756
0.83628 | -33.9992 | -1.3064
0.84239 | -33.9991 | -1.3349
0.84849 | -33.9991 | -1.3634
0.8546 | -33.9996 | -1.3914
0.8607 | -33.9996 | -1.4211
0.86681 | -33.9998 | -1.4498
0.87291 | -33.9993 | -1.4757
0.87901 | -33.9988 | -1.4983
0.88512 | -33.9986 | -1.5215
0.89122 | -33.9985 | -1.5451
0.89733 | -33.9982 | -1.5644
0.90343 | -33.9979 | -1.5858
0.90953 | -33.9981 | -1.6006
0.91564 | -33.9975 | -1.6166
0.92174 | -33.9971 | -1.6308
0.92785 | -33.9966 | -1.6442
0.93395 | -33.9966 | -1.6558
0.94006 | -33.9966 | -1.6692
0.94616 | -33.9964 | -1.6735
0.95226 | -33.9965 | -1.6849
0.95837 | -33.9966 | -1.6964
0.96447 | -33.997 | -1.7064
0.97058 | -33.997 | -1.7161
0.97668 | -33.9968 | -1.7239
0.98279 | -33.9968 | -1.7272
0.98889 | -33.997 | -1.7316
0.99499 | -33.9971 | -1.74
1.0011 | -33.9974 | -1.7445
1.0072 | -33.9972 | -1.7479
1.0133 | -33.9972 | -1.7514
1.0194 | -33.9971 | -1.7566
1.0255 | -33.9968 | -1.7633
1.0316 | -33.9961 | -1.767
1.0377 | -33.9956 | -1.7727
1.0438 | -33.9948 | -1.7768
1.0499 | -33.9932 | -1.7831
1.056 | -33.992 | -1.7876
1.0621 | -33.9908 | -1.7932
1.0682 | -33.9903 | -1.7978
1.0743 | -33.9891 | -1.8054
1.0805 | -33.9885 | -1.812
1.0866 | -33.9867 | -1.8226
1.0927 | -33.9853 | -1.8298
1.0988 | -33.9845 | -1.8402
1.1049 | -33.9829 | -1.8475
1.111 | -33.9813 | -1.8585
1.1171 | -33.9798 | -1.8676
1.1232 | -33.9784 | -1.8817
1.1293 | -33.9772 | -1.8935
1.1354 | -33.9762 | -1.9057
1.1415 | -33.9749 | -1.9197
1.1476 | -33.9735 | -1.9336
1.1537 | -33.9721 | -1.9467
1.1598 | -33.9719 | -1.965
1.1659 | -33.9714 | -1.9769
1.172 | -33.9706 | -1.99
1.1781 | -33.9698 | -2.0038
1.1842 | -33.9692 | -2.0193
1.1903 | -33.97 | -2.0343
1.1964 | -33.9698 | -2.0486
1.2025 | -33.9698 | -2.0638
1.2086 | -33.9702 | -2.0794
1.2147 | -33.9705 | -2.0943
1.2209 | -33.9712 | -2.1075
1.227 | -33.9732 | -2.1224
1.2331 | -33.9742 | -2.1366
1.2392 | -33.9752 | -2.1523
1.2453 | -33.9764 | -2.1641
1.2514 | -33.9778 | -2.1759
1.2575 | -33.9794 | -2.1898
1.2636 | -33.9818 | -2.201
1.2697 | -33.9836 | -2.2125
1.2758 | -33.9853 | -2.2265
1.2819 | -33.9871 | -2.2398
1.288 | -33.9889 | -2.2534
1.2941 | -33.991 | -2.2638
1.3002 | -33.9931 | -2.2718
1.3063 | -33.9948 | -2.2815
1.3124 | -33.9967 | -2.2913
1.3185 | -33.9988 | -2.3017
1.3246 | -34.0007 | -2.3122
1.3307 | -34.0029 | -2.3214
1.3368 | -34.0048 | -2.3287
1.3429 | -34.0065 | -2.3372
1.349 | -34.0081 | -2.3454
1.3551 | -34.0099 | -2.3561
1.3613 | -34.0115 | -2.3637
1.3674 | -34.0128 | -2.3719
1.3735 | -34.0142 | -2.3797
1.3796 | -34.0155 | -2.3868
1.3857 | -34.0166 | -2.3935
1.3918 | -34.0177 | -2.403
1.3979 | -34.019 | -2.4112
1.404 | -34.0197 | -2.421
1.4101 | -34.0207 | -2.4298
1.4162 | -34.0216 | -2.4382
1.4223 | -34.0224 | -2.4483
1.4284 | -34.0233 | -2.4564
1.4345 | -34.0239 | -2.4665
1.4406 | -34.0244 | -2.4777
1.4467 | -34.0251 | -2.4869
1.4528 | -34.0258 | -2.496
1.4589 | -34.0267 | -2.5064
1.465 | -34.0269 | -2.5139
1.4711 | -34.0279 | -2.5228
1.4772 | -34.0284 | -2.535
1.4833 | -34.0291 | -2.5472
1.4894 | -34.0298 | -2.5585
1.4955 | -34.0306 | -2.57
1.5016 | -34.0312 | -2.582
1.5078 | -34.0317 | -2.5966
1.5139 | -34.0324 | -2.6093
1.52 | -34.0329 | -2.6225
1.5261 | -34.0333 | -2.6341
1.5322 | -34.0335 | -2.6468
1.5383 | -34.0338 | -2.659
1.5444 | -34.0345 | -2.6725
1.5505 | -34.0349 | -2.6858
1.5566 | -34.0355 | -2.6986
1.5627 | -34.0361 | -2.7121
1.5688 | -34.0369 | -2.7253
1.5749 | -34.0371 | -2.7377
1.581 | -34.0375 | -2.7545
1.5871 | -34.0381 | -2.7702
1.5932 | -34.0385 | -2.7832
1.5993 | -34.0389 | -2.8001
1.6054 | -34.0392 | -2.8152
1.6115 | -34.0396 | -2.8323
1.6176 | -34.04 | -2.8488
1.6237 | -34.0402 | -2.8634
1.6298 | -34.0405 | -2.8789
1.6359 | -34.0409 | -2.8935
1.642 | -34.0414 | -2.9081
1.6482 | -34.0418 | -2.9253
1.6543 | -34.0423 | -2.9422
1.6604 | -34.0427 | -2.957
1.6665 | -34.0428 | -2.972
1.6726 | -34.0431 | -2.9884
1.6787 | -34.0436 | -3.0055
1.6848 | -34.0435 | -3.022
1.6909 | -34.0438 | -3.0384
1.697 | -34.0441 | -3.0547
1.7031 | -34.0441 | -3.0705
1.7092 | -34.0441 | -3.0855
1.7153 | -34.0443 | -3.1018
1.7214 | -34.0439 | -3.1172
1.7275 | -34.0437 | -3.1331
1.7336 | -34.0435 | -3.1495
1.7397 | -34.0435 | -3.1648
1.7458 | -34.0436 | -3.1815
1.7519 | -34.0434 | -3.1977
1.758 | -34.0433 | -3.2126
1.7641 | -34.0434 | -3.2271
1.7702 | -34.0432 | -3.243
1.7763 | -34.0432 | -3.2567
1.7824 | -34.043 | -3.2729
1.7885 | -34.0428 | -3.287
1.7947 | -34.0429 | -3.3019
1.8008 | -34.0428 | -3.3147
1.8069 | -34.0427 | -3.3283
1.813 | -34.0427 | -3.3425
1.8191 | -34.0424 | -3.3568
1.8252 | -34.0422 | -3.3707
1.8313 | -34.0423 | -3.3851
1.8374 | -34.0423 | -3.3986
1.8435 | -34.0423 | -3.4108
1.8496 | -34.042 | -3.4253
1.8557 | -34.042 | -3.4401
1.8618 | -34.0419 | -3.4518
1.8679 | -34.0419 | -3.4648
1.874 | -34.0419 | -3.4779
1.8801 | -34.0418 | -3.4909
1.8862 | -34.042 | -3.5036
1.8923 | -34.0419 | -3.5157
1.8984 | -34.0416 | -3.5263
1.9045 | -34.0415 | -3.5373
1.9106 | -34.0412 | -3.5491
1.9167 | -34.0409 | -3.5603
1.9228 | -34.0409 | -3.5711
1.9289 | -34.0405 | -3.5799
1.9351 | -34.0406 | -3.5888
1.9412 | -34.0407 | -3.599
1.9473 | -34.0408 | -3.6109
1.9534 | -34.0407 | -3.622
1.9595 | -34.0405 | -3.6334
1.9656 | -34.04 | -3.6441
1.9717 | -34.0398 | -3.6545
1.9778 | -34.0398 | -3.666
1.9839 | -34.0399 | -3.6773
1.99 | -34.0399 | -3.6871
1.9961 | -34.0396 | -3.6975
2.0022 | -34.0396 | -3.7074
2.0083 | -34.0392 | -3.7172
2.0144 | -34.0392 | -3.727
2.0205 | -34.0393 | -3.7388
2.0266 | -34.0394 | -3.7503
2.0327 | -34.0395 | -3.7614
2.0388 | -34.0397 | -3.7753
2.0449 | -34.0398 | -3.7855
2.051 | -34.0399 | -3.7955
2.0571 | -34.0397 | -3.8068
2.0632 | -34.0396 | -3.8173
2.0693 | -34.0396 | -3.8284
2.0754 | -34.0399 | -3.8402
2.0816 | -34.04 | -3.8522
2.0877 | -34.0399 | -3.8633
2.0938 | -34.0402 | -3.8756
2.0999 | -34.0402 | -3.8872
2.106 | -34.0404 | -3.8992
2.1121 | -34.0404 | -3.9115
2.1182 | -34.0406 | -3.9215
2.1243 | -34.0408 | -3.932
2.1304 | -34.0409 | -3.9422
2.1365 | -34.0412 | -3.9538
2.1426 | -34.0414 | -3.9672
2.1487 | -34.0418 | -3.9788
2.1548 | -34.0423 | -3.9919
2.1609 | -34.0429 | -4.0039
2.167 | -34.0434 | -4.0165
2.1731 | -34.0437 | -4.0284
2.1792 | -34.0444 | -4.0396
2.1853 | -34.0449 | -4.0525
2.1914 | -34.0454 | -4.0647
2.1975 | -34.046 | -4.0769
2.2036 | -34.0465 | -4.0902
2.2097 | -34.0469 | -4.1021
2.2158 | -34.0479 | -4.1123
2.222 | -34.0485 | -4.1224
2.2281 | -34.0488 | -4.1332
2.2342 | -34.0493 | -4.1464
2.2403 | -34.0498 | -4.1564
2.2464 | -34.0506 | -4.1674
2.2525 | -34.051 | -4.1782
2.2586 | -34.052 | -4.188
2.2647 | -34.0528 | -4.1973
2.2708 | -34.0534 | -4.209
2.2769 | -34.0543 | -4.2204
2.283 | -34.0548 | -4.2297
2.2891 | -34.0553 | -4.2401
2.2952 | -34.0559 | -4.2498
2.3013 | -34.0566 | -4.2599
2.3074 | -34.0573 | -4.2694
2.3135 | -34.0578 | -4.2808
2.3196 | -34.0584 | -4.2892
2.3257 | -34.0588 | -4.2997
2.3318 | -34.0591 | -4.3086
2.3379 | -34.0596 | -4.32
2.344 | -34.0599 | -4.3308
2.3501 | -34.0604 | -4.3415
2.3562 | -34.0607 | -4.3497
2.3623 | -34.061 | -4.3583
2.3685 | -34.0613 | -4.3672
2.3746 | -34.0615 | -4.3776
2.3807 | -34.062 | -4.3865
2.3868 | -34.0622 | -4.3954
2.3929 | -34.0626 | -4.4046
2.399 | -34.0629 | -4.4125
2.4051 | -34.0634 | -4.4219
2.4112 | -34.0637 | -4.4308
2.4173 | -34.064 | -4.4396
2.4234 | -34.064 | -4.448
2.4295 | -34.0643 | -4.4565
2.4356 | -34.0644 | -4.4657
2.4417 | -34.0644 | -4.4757
2.4478 | -34.0645 | -4.4851
2.4539 | -34.0648 | -4.4929
2.46 | -34.0645 | -4.5025
2.4661 | -34.0645 | -4.5117
2.4722 | -34.0648 | -4.5202
2.4783 | -34.0649 | -4.5285
2.4844 | -34.0649 | -4.5369
2.4905 | -34.0648 | -4.546
2.4966 | -34.0645 | -4.5548
2.5027 | -34.0645 | -4.5627
2.5089 | -34.0644 | -4.5711
2.515 | -34.0641 | -4.5803
2.5211 | -34.0641 | -4.5885
2.5272 | -34.0638 | -4.5974
2.5333 | -34.0636 | -4.6068
2.5394 | -34.0635 | -4.6158
2.5455 | -34.0633 | -4.6251
2.5516 | -34.0633 | -4.6337
2.5577 | -34.0632 | -4.6436
2.5638 | -34.063 | -4.6539
2.5699 | -34.0627 | -4.6641
2.576 | -34.0625 | -4.6745
2.5821 | -34.0623 | -4.6843
2.5882 | -34.0621 | -4.6938
2.5943 | -34.0619 | -4.7035
2.6004 | -34.0618 | -4.7135
2.6065 | -34.0616 | -4.7218
2.6126 | -34.0615 | -4.7308
2.6187 | -34.0614 | -4.7395
2.6248 | -34.0612 | -4.7484
2.6309 | -34.061 | -4.7597
2.637 | -34.0611 | -4.7709
2.6431 | -34.061 | -4.7803
2.6492 | -34.0609 | -4.7903
2.6554 | -34.0609 | -4.8004
2.6615 | -34.0607 | -4.8103
2.6676 | -34.0606 | -4.8207
2.6737 | -34.0608 | -4.8314
2.6798 | -34.0608 | -4.8413
2.6859 | -34.0608 | -4.8519
2.692 | -34.0607 | -4.8615
2.6981 | -34.0606 | -4.8727
2.7042 | -34.0607 | -4.8848
2.7103 | -34.0607 | -4.8948
2.7164 | -34.0609 | -4.9055
2.7225 | -34.061 | -4.9168
2.7286 | -34.0611 | -4.928
2.7347 | -34.0611 | -4.938
2.7408 | -34.0614 | -4.9472
2.7469 | -34.0614 | -4.9585
2.753 | -34.0617 | -4.9692
2.7591 | -34.0618 | -4.9811
2.7652 | -34.0618 | -4.9927
2.7713 | -34.0618 | -5.0025
2.7774 | -34.0623 | -5.014
2.7835 | -34.0624 | -5.0246
2.7896 | -34.0625 | -5.0339
2.7958 | -34.0629 | -5.0441
2.8019 | -34.0635 | -5.0539
2.808 | -34.0638 | -5.0658
2.8141 | -34.0642 | -5.0788
2.8202 | -34.0646 | -5.091
2.8263 | -34.065 | -5.102
2.8324 | -34.0653 | -5.1145
2.8385 | -34.0657 | -5.1266
2.8446 | -34.0663 | -5.1383
2.8507 | -34.0667 | -5.1505
2.8568 | -34.0672 | -5.1616
2.8629 | -34.0676 | -5.1729
2.869 | -34.0681 | -5.1843
2.8751 | -34.0685 | -5.1963
2.8812 | -34.0692 | -5.2096
2.8873 | -34.0695 | -5.221
2.8934 | -34.0699 | -5.2329
2.8995 | -34.0703 | -5.2449
2.9056 | -34.0706 | -5.2557
2.9117 | -34.0711 | -5.2673
2.9178 | -34.0718 | -5.28
2.9239 | -34.0722 | -5.2935
2.93 | -34.0727 | -5.307
2.9361 | -34.0733 | -5.3184
2.9423 | -34.0739 | -5.3305
2.9484 | -34.0744 | -5.344
2.9545 | -34.0751 | -5.3574
2.9606 | -34.0756 | -5.3708
2.9667 | -34.0762 | -5.3827
2.9728 | -34.0766 | -5.395
2.9789 | -34.0771 | -5.4078
2.985 | -34.0775 | -5.4216
2.9911 | -34.078 | -5.4342
2.9972 | -34.0786 | -5.447
3.0033 | -34.0792 | -5.4601
3.0094 | -34.0798 | -5.4732
3.0155 | -34.0804 | -5.4858
3.0216 | -34.0809 | -5.4993
3.0277 | -34.0814 | -5.5131
3.0338 | -34.082 | -5.5267
3.0399 | -34.0826 | -5.5406
3.046 | -34.0829 | -5.5536
3.0521 | -34.0835 | -5.5661
3.0582 | -34.0838 | -5.5801
3.0643 | -34.0844 | -5.5932
3.0704 | -34.0847 | -5.6067
3.0765 | -34.085 | -5.6189
3.0827 | -34.0854 | -5.6321
3.0888 | -34.0857 | -5.645
3.0949 | -34.0861 | -5.6581
3.101 | -34.0866 | -5.6723
3.1071 | -34.0869 | -5.6865
3.1132 | -34.0873 | -5.7006
3.1193 | -34.0878 | -5.7137
3.1254 | -34.0882 | -5.726
3.1315 | -34.0886 | -5.7404
3.1376 | -34.0889 | -5.7533
3.1437 | -34.0892 | -5.767
3.1498 | -34.0896 | -5.7813
3.1559 | -34.0899 | -5.7946
3.162 | -34.09 | -5.8087
3.1681 | -34.0901 | -5.823
3.1742 | -34.0905 | -5.8368
3.1803 | -34.0908 | -5.8501
3.1864 | -34.0912 | -5.8624
3.1925 | -34.0915 | -5.8761
3.1986 | -34.0916 | -5.889
3.2047 | -34.0917 | -5.9031
3.2108 | -34.092 | -5.9148
3.2169 | -34.0924 | -5.9287
3.223 | -34.0924 | -5.9421
3.2292 | -34.0924 | -5.9561
3.2353 | -34.0923 | -5.9657
3.2414 | -34.0924 | -5.9787
3.2475 | -34.0924 | -5.9898
3.2536 | -34.0927 | -6.0036
3.2597 | -34.0929 | -6.0148
3.2658 | -34.0931 | -6.0255
3.2719 | -34.0932 | -6.0366
3.278 | -34.0933 | -6.0488
3.2841 | -34.0935 | -6.0607
3.2902 | -34.0936 | -6.0733
3.2963 | -34.0937 | -6.085
3.3024 | -34.0936 | -6.0952
3.3085 | -34.0938 | -6.1078
3.3146 | -34.0937 | -6.1187
3.3207 | -34.0939 | -6.1309
3.3268 | -34.0941 | -6.1431
3.3329 | -34.0942 | -6.1529
3.339 | -34.0943 | -6.1625
3.3451 | -34.0943 | -6.1739
3.3512 | -34.0944 | -6.1833
3.3573 | -34.0946 | -6.1947
3.3634 | -34.0947 | -6.2062
3.3696 | -34.0946 | -6.2181
3.3757 | -34.0946 | -6.2279
3.3818 | -34.0949 | -6.2396
3.3879 | -34.0948 | -6.2499
3.394 | -34.0948 | -6.2604
3.4001 | -34.0948 | -6.2712
3.4062 | -34.0948 | -6.2814
3.4123 | -34.0949 | -6.2911
3.4184 | -34.0949 | -6.3021
3.4245 | -34.095 | -6.3126
3.4306 | -34.0949 | -6.3223
3.4367 | -34.0951 | -6.332
3.4428 | -34.0954 | -6.3411
3.4489 | -34.0953 | -6.3522
3.455 | -34.0952 | -6.3626
3.4611 | -34.0953 | -6.3737
3.4672 | -34.0953 | -6.385
3.4733 | -34.0953 | -6.3951
3.4794 | -34.0953 | -6.4055
3.4855 | -34.0956 | -6.415
3.4916 | -34.0957 | -6.4247
3.4977 | -34.0956 | -6.4356
3.5038 | -34.0958 | -6.4455
3.5099 | -34.0958 | -6.455
3.5161 | -34.096 | -6.4644
3.5222 | -34.0961 | -6.4744
3.5283 | -34.096 | -6.4844
3.5344 | -34.096 | -6.4951
3.5405 | -34.0962 | -6.505
3.5466 | -34.0963 | -6.5147
3.5527 | -34.0964 | -6.5253
3.5588 | -34.0965 | -6.535
3.5649 | -34.0965 | -6.545
3.571 | -34.0966 | -6.5551
3.5771 | -34.0967 | -6.5658
3.5832 | -34.0968 | -6.5753
3.5893 | -34.097 | -6.5851
3.5954 | -34.0971 | -6.5955
3.6015 | -34.0973 | -6.6058
3.6076 | -34.0975 | -6.6154
3.6137 | -34.0976 | -6.6251
3.6198 | -34.0979 | -6.634
3.6259 | -34.098 | -6.6447
3.632 | -34.0982 | -6.6552
3.6381 | -34.0984 | -6.665
3.6442 | -34.0987 | -6.6754
3.6503 | -34.0989 | -6.6843
3.6565 | -34.0992 | -6.696
3.6626 | -34.0993 | -6.7066
3.6687 | -34.0996 | -6.7172
3.6748 | -34.0999 | -6.7269
3.6809 | -34.1003 | -6.7374
3.687 | -34.1005 | -6.7485
3.6931 | -34.1008 | -6.7584
3.6992 | -34.1011 | -6.7689
3.7053 | -34.1013 | -6.7789
3.7114 | -34.1015 | -6.7901
3.7175 | -34.1018 | -6.7999
3.7236 | -34.102 | -6.8085
3.7297 | -34.1022 | -6.8192
3.7358 | -34.1024 | -6.8309
3.7419 | -34.1026 | -6.8424
3.748 | -34.1028 | -6.853
3.7541 | -34.1029 | -6.8639
3.7602 | -34.1033 | -6.8747
3.7663 | -34.1035 | -6.8848
3.7724 | -34.1038 | -6.8957
3.7785 | -34.1042 | -6.906
3.7846 | -34.1044 | -6.9166
3.7907 | -34.1048 | -6.9275
3.7969 | -34.105 | -6.9389
3.803 | -34.1054 | -6.9489
3.8091 | -34.1057 | -6.9591
3.8152 | -34.106 | -6.9703
3.8213 | -34.1062 | -6.9818
3.8274 | -34.1066 | -6.9922
3.8335 | -34.107 | -7.0037
3.8396 | -34.1075 | -7.0151
3.8457 | -34.1078 | -7.0253
3.8518 | -34.1083 | -7.0361
3.8579 | -34.1087 | -7.0466
3.864 | -34.1091 | -7.0573
3.8701 | -34.1095 | -7.0682
3.8762 | -34.1098 | -7.0793
3.8823 | -34.1101 | -7.0898
3.8884 | -34.1105 | -7.1001
3.8945 | -34.111 | -7.1114
3.9006 | -34.1114 | -7.1229
3.9067 | -34.1118 | -7.1347
3.9128 | -34.1122 | -7.1459
3.9189 | -34.1126 | -7.1565
3.925 | -34.1128 | -7.1673
3.9311 | -34.1133 | -7.1772
3.9372 | -34.1137 | -7.1887
3.9434 | -34.114 | -7.1991
3.9495 | -34.1144 | -7.2106
3.9556 | -34.1149 | -7.2218
3.9617 | -34.1153 | -7.2326
3.9678 | -34.1156 | -7.2432
3.9739 | -34.1161 | -7.2544
3.98 | -34.1164 | -7.2662
3.9861 | -34.1167 | -7.2766
3.9922 | -34.117 | -7.2882
3.9983 | -34.1174 | -7.2981
4.0044 | -34.1177 | -7.3077
4.0105 | -34.1181 | -7.3185
4.0166 | -34.1186 | -7.3295
4.0227 | -34.119 | -7.3398
4.0288 | -34.1192 | -7.3511
4.0349 | -34.1196 | -7.3609
4.041 | -34.1199 | -7.3718
4.0471 | -34.1202 | -7.382
4.0532 | -34.1207 | -7.3931
4.0593 | -34.1209 | -7.4046
4.0654 | -34.1213 | -7.4153
4.0715 | -34.1216 | -7.4266
4.0776 | -34.122 | -7.4369
4.0838 | -34.1224 | -7.4483
4.0899 | -34.1227 | -7.4595
4.096 | -34.123 | -7.4698
4.1021 | -34.1234 | -7.4805
4.1082 | -34.1238 | -7.4907
4.1143 | -34.1241 | -7.5006
4.1204 | -34.1245 | -7.5113
4.1265 | -34.1249 | -7.5216
4.1326 | -34.1253 | -7.5318
4.1387 | -34.1258 | -7.543
4.1448 | -34.1261 | -7.5539
4.1509 | -34.1263 | -7.5645
4.157 | -34.1266 | -7.5757
4.1631 | -34.1269 | -7.5869
4.1692 | -34.1273 | -7.5975
4.1753 | -34.1277 | -7.6083
4.1814 | -34.1279 | -7.618
4.1875 | -34.1283 | -7.6294
4.1936 | -34.1285 | -7.6406
4.1997 | -34.1289 | -7.6506
4.2058 | -34.1291 | -7.6621
4.2119 | -34.1294 | -7.6718
4.218 | -34.1297 | -7.6818
4.2241 | -34.1301 | -7.6913
4.2303 | -34.1305 | -7.7025
4.2364 | -34.1308 | -7.7138
4.2425 | -34.1311 | -7.7243
4.2486 | -34.1314 | -7.7354
4.2547 | -34.1315 | -7.7445
4.2608 | -34.1318 | -7.7544
4.2669 | -34.1321 | -7.7647
4.273 | -34.1324 | -7.7754
4.2791 | -34.1326 | -7.786
4.2852 | -34.1328 | -7.7962
4.2913 | -34.1329 | -7.8054
4.2974 | -34.1332 | -7.8163
4.3035 | -34.1334 | -7.8271
4.3096 | -34.1336 | -7.8374
4.3157 | -34.1338 | -7.8473
4.3218 | -34.1341 | -7.8572
4.3279 | -34.1343 | -7.8675
4.334 | -34.1345 | -7.8773
4.3401 | -34.1347 | -7.8871
4.3462 | -34.135 | -7.8978
4.3523 | -34.1354 | -7.9074
4.3584 | -34.1355 | -7.9158
4.3645 | -34.1357 | -7.9253
4.3707 | -34.136 | -7.9345
4.3768 | -34.1362 | -7.9448
4.3829 | -34.1363 | -7.955
4.389 | -34.1366 | -7.9647
4.3951 | -34.1368 | -7.9741
4.4012 | -34.137 | -7.9838
4.4073 | -34.1373 | -7.9938
4.4134 | -34.1375 | -8.0039
4.4195 | -34.1377 | -8.0134
4.4256 | -34.138 | -8.0229
4.4317 | -34.1383 | -8.0324
4.4378 | -34.1384 | -8.0418
4.4439 | -34.1386 | -8.0511
4.45 | -34.1387 | -8.061
4.4561 | -34.1388 | -8.0707
4.4622 | -34.139 | -8.0803
4.4683 | -34.1391 | -8.0894
4.4744 | -34.1393 | -8.0995
4.4805 | -34.1396 | -8.1084
4.4866 | -34.1398 | -8.1176
4.4927 | -34.14 | -8.1277
4.4988 | -34.1402 | -8.137
4.5049 | -34.1404 | -8.146
4.511 | -34.1407 | -8.155
4.5172 | -34.141 | -8.1647
4.5233 | -34.1412 | -8.1752
4.5294 | -34.1414 | -8.1845
4.5355 | -34.1416 | -8.1938
4.5416 | -34.1419 | -8.2028
4.5477 | -34.1421 | -8.212
4.5538 | -34.1423 | -8.2206
4.5599 | -34.1426 | -8.23
4.566 | -34.1428 | -8.2396
4.5721 | -34.143 | -8.2487
4.5782 | -34.1433 | -8.2585
4.5843 | -34.1435 | -8.268
4.5904 | -34.1438 | -8.2772
4.5965 | -34.1441 | -8.2858
4.6026 | -34.1444 | -8.2964
4.6087 | -34.1446 | -8.3053
4.6148 | -34.1448 | -8.3159
4.6209 | -34.1452 | -8.3243
4.627 | -34.1455 | -8.3342
4.6331 | -34.1456 | -8.343
4.6392 | -34.1458 | -8.3522
4.6453 | -34.1461 | -8.3615
4.6514 | -34.1464 | -8.3709
4.6576 | -34.1464 | -8.3813
4.6637 | -34.1466 | -8.3889
4.6698 | -34.1469 | -8.3974
4.6759 | -34.147 | -8.4071
4.682 | -34.1471 | -8.4161
4.6881 | -34.1473 | -8.4249
4.6942 | -34.1476 | -8.434
4.7003 | -34.1478 | -8.4434
4.7064 | -34.148 | -8.4526
4.7125 | -34.1484 | -8.462
4.7186 | -34.1485 | -8.4703
4.7247 | -34.1489 | -8.4799
4.7308 | -34.1492 | -8.4891
4.7369 | -34.1493 | -8.4986
4.743 | -34.1495 | -8.5089
4.7491 | -34.1497 | -8.5185
4.7552 | -34.1498 | -8.5276
4.7613 | -34.15 | -8.5371
4.7674 | -34.1503 | -8.5456
4.7735 | -34.1504 | -8.5557
4.7796 | -34.1507 | -8.5654
4.7857 | -34.1508 | -8.575
4.7918 | -34.151 | -8.5835
4.7979 | -34.1513 | -8.5925
4.8041 | -34.1514 | -8.6017
4.8102 | -34.1518 | -8.6107
4.8163 | -34.152 | -8.6201
4.8224 | -34.1521 | -8.6287
4.8285 | -34.1525 | -8.6376
4.8346 | -34.1527 | -8.6471
4.8407 | -34.1529 | -8.6563
4.8468 | -34.1531 | -8.666
4.8529 | -34.1533 | -8.6757
4.859 | -34.1536 | -8.6852
4.8651 | -34.1538 | -8.6944
4.8712 | -34.154 | -8.7037
4.8773 | -34.1542 | -8.7132
4.8834 | -34.1545 | -8.7229
4.8895 | -34.1548 | -8.7328
4.8956 | -34.1549 | -8.7427
4.9017 | -34.1551 | -8.7521
4.9078 | -34.1553 | -8.762
4.9139 | -34.1555 | -8.7704
4.92 | -34.1558 | -8.7794
4.9261 | -34.1561 | -8.7881
4.9322 | -34.1565 | -8.7972
4.9383 | -34.1568 | -8.8062
4.9445 | -34.1571 | -8.8162
4.9506 | -34.1573 | -8.825
4.9567 | -34.1576 | -8.8342
4.9628 | -34.1578 | -8.8435
4.9689 | -34.158 | -8.8532
4.975 | -34.1582 | -8.8625
4.9811 | -34.1585 | -8.8722
4.9872 | -34.1587 | -8.882
4.9933 | -34.1591 | -8.8916
4.9994 | -34.1595 | -8.9014
5.0055 | -34.1598 | -8.911
5.0116 | -34.1601 | -8.9212
5.0177 | -34.1605 | -8.9313
5.0238 | -34.1608 | -8.9418
5.0299 | -34.1613 | -8.9527
5.036 | -34.1615 | -8.9622
5.0421 | -34.1617 | -8.9718
5.0482 | -34.162 | -8.9821
5.0543 | -34.1623 | -8.9913
5.0604 | -34.1627 | -9.0015
5.0665 | -34.1631 | -9.0111
5.0726 | -34.1634 | -9.0215
5.0787 | -34.1637 | -9.0317
5.0848 | -34.164 | -9.0417
5.091 | -34.1643 | -9.0516
5.0971 | -34.1646 | -9.0616
5.1032 | -34.165 | -9.0719
5.1093 | -34.1653 | -9.0818
5.1154 | -34.1656 | -9.0918
5.1215 | -34.1658 | -9.1017
5.1276 | -34.1661 | -9.1115
5.1337 | -34.1666 | -9.1211
5.1398 | -34.1668 | -9.1309
5.1459 | -34.1672 | -9.1413
5.152 | -34.1675 | -9.1523
5.1581 | -34.1677 | -9.1629
5.1642 | -34.168 | -9.1729
5.1703 | -34.1684 | -9.1828
5.1764 | -34.1687 | -9.1927
5.1825 | -34.169 | -9.203
5.1886 | -34.1693 | -9.2127
5.1947 | -34.1696 | -9.2235
5.2008 | -34.17 | -9.2333
5.2069 | -34.1704 | -9.2445
5.213 | -34.1707 | -9.254
5.2191 | -34.1711 | -9.2643
5.2252 | -34.1714 | -9.2745
5.2314 | -34.1716 | -9.2853
5.2375 | -34.172 | -9.2951
5.2436 | -34.1723 | -9.3053
5.2497 | -34.1727 | -9.3151
5.2558 | -34.173 | -9.3261
5.2619 | -34.1733 | -9.3369
5.268 | -34.1737 | -9.3474
5.2741 | -34.1741 | -9.3586
5.2802 | -34.1744 | -9.3694
5.2863 | -34.1748 | -9.3796
5.2924 | -34.1751 | -9.3907
5.2985 | -34.1755 | -9.4021
5.3046 | -34.176 | -9.4133
5.3107 | -34.1763 | -9.4241
5.3168 | -34.1768 | -9.4345
5.3229 | -34.1771 | -9.4444
5.329 | -34.1775 | -9.4543
5.3351 | -34.1779 | -9.4644
5.3412 | -34.1783 | -9.4746
5.3473 | -34.1787 | -9.4854
5.3534 | -34.1791 | -9.4963
5.3595 | -34.1794 | -9.507
5.3656 | -34.1798 | -9.518
5.3717 | -34.1801 | -9.5285
5.3779 | -34.1805 | -9.5394
5.384 | -34.1808 | -9.5502
5.3901 | -34.1812 | -9.5609
5.3962 | -34.1815 | -9.5716
5.4023 | -34.1819 | -9.5824
5.4084 | -34.1823 | -9.5928
5.4145 | -34.1826 | -9.6038
5.4206 | -34.1829 | -9.6151
5.4267 | -34.1833 | -9.6264
5.4328 | -34.1837 | -9.6366
5.4389 | -34.184 | -9.6468
5.445 | -34.1843 | -9.6581
5.4511 | -34.1847 | -9.6685
5.4572 | -34.1851 | -9.6797
5.4633 | -34.1855 | -9.69
5.4694 | -34.1857 | -9.7005
5.4755 | -34.1861 | -9.7114
5.4816 | -34.1865 | -9.722
5.4877 | -34.1868 | -9.7327
5.4938 | -34.1872 | -9.7431
5.4999 | -34.1875 | -9.7538
5.506 | -34.1879 | -9.7645
5.5121 | -34.1883 | -9.7756
5.5183 | -34.1886 | -9.7856
5.5244 | -34.1888 | -9.7964
5.5305 | -34.1892 | -9.8073
5.5366 | -34.1896 | -9.8177
5.5427 | -34.1899 | -9.829
5.5488 | -34.1902 | -9.84
5.5549 | -34.1905 | -9.8507
5.561 | -34.1909 | -9.862
5.5671 | -34.1913 | -9.8724
5.5732 | -34.1915 | -9.8833
5.5793 | -34.1919 | -9.8949
5.5854 | -34.1924 | -9.9057
5.5915 | -34.1926 | -9.9167
5.5976 | -34.1929 | -9.9277
5.6037 | -34.1933 | -9.9384
5.6098 | -34.1936 | -9.9489
5.6159 | -34.1939 | -9.9599
5.622 | -34.1943 | -9.9703
5.6281 | -34.1946 | -9.9812
5.6342 | -34.195 | -9.9921
5.6403 | -34.1953 | -10.0023
5.6464 | -34.1957 | -10.0131
5.6525 | -34.1959 | -10.0236
5.6586 | -34.1964 | -10.0335
5.6648 | -34.1967 | -10.0448
5.6709 | -34.197 | -10.056
5.677 | -34.1974 | -10.0675
5.6831 | -34.1977 | -10.0777
5.6892 | -34.198 | -10.0885
5.6953 | -34.1983 | -10.0991
5.7014 | -34.1987 | -10.1104
5.7075 | -34.199 | -10.1214
5.7136 | -34.1993 | -10.1319
5.7197 | -34.1997 | -10.1425
5.7258 | -34.1999 | -10.153
5.7319 | -34.2003 | -10.1638
5.738 | -34.2006 | -10.1745
5.7441 | -34.201 | -10.1852
5.7502 | -34.2014 | -10.1955
5.7563 | -34.2017 | -10.2059
5.7624 | -34.202 | -10.2174
5.7685 | -34.2023 | -10.2279
5.7746 | -34.2025 | -10.2384
5.7807 | -34.2028 | -10.2488
5.7868 | -34.2031 | -10.2594
5.7929 | -34.2034 | -10.27
5.799 | -34.2038 | -10.2805
5.8052 | -34.2041 | -10.2913
5.8113 | -34.2043 | -10.302
5.8174 | -34.2046 | -10.3127
5.8235 | -34.2049 | -10.3232
5.8296 | -34.2052 | -10.3345
5.8357 | -34.2054 | -10.3451
5.8418 | -34.2058 | -10.356
5.8479 | -34.206 | -10.3669
5.854 | -34.2063 | -10.3779
5.8601 | -34.2066 | -10.3883
5.8662 | -34.2069 | -10.3993
5.8723 | -34.2072 | -10.4093
5.8784 | -34.2075 | -10.4198
5.8845 | -34.2078 | -10.4306
5.8906 | -34.2081 | -10.4412
5.8967 | -34.2084 | -10.4514
5.9028 | -34.2088 | -10.4621
5.9089 | -34.2091 | -10.4721
5.915 | -34.2094 | -10.4826
5.9211 | -34.2097 | -10.4933
5.9272 | -34.21 | -10.5038
5.9333 | -34.2103 | -10.5137
5.9394 | -34.2106 | -10.524
5.9455 | -34.211 | -10.5342
5.9517 | -34.2113 | -10.5447
5.9578 | -34.2117 | -10.5556
5.9639 | -34.212 | -10.5658
5.97 | -34.2123 | -10.5764
5.9761 | -34.2127 | -10.5866
5.9822 | -34.213 | -10.5973
5.9883 | -34.2134 | -10.6076
5.9944 | -34.2137 | -10.6182
6.0005 | -34.2141 | -10.629
6.0066 | -34.2143 | -10.6391
6.0127 | -34.2147 | -10.6488
6.0188 | -34.215 | -10.6597
6.0249 | -34.2153 | -10.6701
6.031 | -34.2156 | -10.6803
6.0371 | -34.216 | -10.6911
6.0432 | -34.2164 | -10.7015
6.0493 | -34.2167 | -10.7117
6.0554 | -34.2171 | -10.7221
6.0615 | -34.2174 | -10.733
6.0676 | -34.2177 | -10.7441
6.0737 | -34.2181 | -10.7548
6.0798 | -34.2183 | -10.7651
6.0859 | -34.2187 | -10.7755
6.0921 | -34.219 | -10.7856
6.0982 | -34.2194 | -10.7963
6.1043 | -34.2198 | -10.8066
6.1104 | -34.2202 | -10.8175
6.1165 | -34.2205 | -10.8272
6.1226 | -34.2209 | -10.837
6.1287 | -34.2212 | -10.8475
6.1348 | -34.2217 | -10.858
6.1409 | -34.222 | -10.8682
6.147 | -34.2223 | -10.8785
6.1531 | -34.2227 | -10.8888
6.1592 | -34.223 | -10.8992
6.1653 | -34.2234 | -10.9093
6.1714 | -34.2237 | -10.9197
6.1775 | -34.2241 | -10.9304
6.1836 | -34.2243 | -10.9405
6.1897 | -34.2247 | -10.951
6.1958 | -34.225 | -10.9616
6.2019 | -34.2254 | -10.9722
6.208 | -34.2257 | -10.983
6.2141 | -34.2261 | -10.9936
6.2202 | -34.2265 | -11.0044
6.2263 | -34.2268 | -11.0146
6.2325 | -34.2272 | -11.0245
6.2386 | -34.2276 | -11.0345
6.2447 | -34.2279 | -11.0451
6.2508 | -34.2282 | -11.0554
6.2569 | -34.2286 | -11.066
6.263 | -34.229 | -11.076
6.2691 | -34.2294 | -11.0863
6.2752 | -34.2297 | -11.0965
6.2813 | -34.2301 | -11.1068
6.2874 | -34.2304 | -11.1172
6.2935 | -34.2308 | -11.1273
6.2996 | -34.2312 | -11.1377
6.3057 | -34.2315 | -11.1479
6.3118 | -34.2319 | -11.1582
6.3179 | -34.2323 | -11.1685
6.324 | -34.2326 | -11.179
6.3301 | -34.233 | -11.1891
6.3362 | -34.2334 | -11.1995
6.3423 | -34.2338 | -11.2099
6.3484 | -34.2342 | -11.2204
6.3545 | -34.2346 | -11.2301
6.3606 | -34.235 | -11.2401
6.3667 | -34.2354 | -11.2501
6.3728 | -34.2358 | -11.2605
6.379 | -34.2361 | -11.2707
6.3851 | -34.2365 | -11.2811
6.3912 | -34.2369 | -11.291
6.3973 | -34.2373 | -11.3011
6.4034 | -34.2376 | -11.3113
6.4095 | -34.238 | -11.3218
6.4156 | -34.2383 | -11.3317
6.4217 | -34.2387 | -11.3419
6.4278 | -34.239 | -11.3521
6.4339 | -34.2395 | -11.3623
6.44 | -34.2399 | -11.3721
6.4461 | -34.2402 | -11.3829
6.4522 | -34.2406 | -11.3925
6.4583 | -34.2409 | -11.4028
6.4644 | -34.2412 | -11.4134
6.4705 | -34.2417 | -11.4238
6.4766 | -34.242 | -11.4344
6.4827 | -34.2424 | -11.4451
6.4888 | -34.2427 | -11.4551
6.4949 | -34.2431 | -11.4649
6.501 | -34.2434 | -11.4755
6.5071 | -34.2438 | -11.4858
6.5132 | -34.2443 | -11.4959
6.5194 | -34.2446 | -11.506
6.5255 | -34.245 | -11.5158
6.5316 | -34.2453 | -11.5262
6.5377 | -34.2457 | -11.5367
6.5438 | -34.2461 | -11.5468
6.5499 | -34.2465 | -11.5568
6.556 | -34.2469 | -11.5668
6.5621 | -34.2472 | -11.5773
6.5682 | -34.2475 | -11.5873
6.5743 | -34.2479 | -11.5983
6.5804 | -34.2482 | -11.6085
6.5865 | -34.2485 | -11.619
6.5926 | -34.249 | -11.6292
6.5987 | -34.2493 | -11.639
6.6048 | -34.2497 | -11.6499
6.6109 | -34.2501 | -11.66
6.617 | -34.2504 | -11.6712
6.6231 | -34.2507 | -11.6816
6.6292 | -34.2512 | -11.6917
6.6353 | -34.2515 | -11.7029
6.6414 | -34.252 | -11.7135
6.6475 | -34.2523 | -11.7239
6.6536 | -34.2527 | -11.7338
6.6597 | -34.2531 | -11.7437
6.6659 | -34.2535 | -11.7538
6.672 | -34.2539 | -11.7641
6.6781 | -34.2542 | -11.7741
6.6842 | -34.2546 | -11.7846
6.6903 | -34.255 | -11.7947
6.6964 | -34.2554 | -11.805
6.7025 | -34.2557 | -11.8152
6.7086 | -34.2561 | -11.8253
6.7147 | -34.2565 | -11.8357
6.7208 | -34.2568 | -11.8464
6.7269 | -34.2571 | -11.8563
6.733 | -34.2574 | -11.8657
6.7391 | -34.2579 | -11.8762
6.7452 | -34.2583 | -11.8858
6.7513 | -34.2587 | -11.8957
6.7574 | -34.2591 | -11.9062
6.7635 | -34.2595 | -11.9158
6.7696 | -34.26 | -11.9261
6.7757 | -34.2604 | -11.9367
6.7818 | -34.2608 | -11.9469
6.7879 | -34.2611 | -11.9572
6.794 | -34.2615 | -11.9672
6.8001 | -34.2618 | -11.9772
6.8063 | -34.2622 | -11.9876
6.8124 | -34.2626 | -11.9979
6.8185 | -34.2629 | -12.0081
6.8246 | -34.2633 | -12.0183
6.8307 | -34.2637 | -12.0283
6.8368 | -34.2641 | -12.0389
6.8429 | -34.2645 | -12.0489
6.849 | -34.2648 | -12.0592
6.8551 | -34.2652 | -12.0694
6.8612 | -34.2656 | -12.0794
6.8673 | -34.266 | -12.0894
6.8734 | -34.2663 | -12.1003
6.8795 | -34.2668 | -12.1108
6.8856 | -34.2671 | -12.1204
6.8917 | -34.2675 | -12.13
6.8978 | -34.2678 | -12.1401
6.9039 | -34.2682 | -12.1501
6.91 | -34.2686 | -12.1604
6.9161 | -34.269 | -12.1703
6.9222 | -34.2693 | -12.1805
6.9283 | -34.2697 | -12.1903
6.9344 | -34.27 | -12.2002
6.9405 | -34.2704 | -12.2101
6.9466 | -34.2708 | -12.2205
6.9528 | -34.2711 | -12.2307
6.9589 | -34.2715 | -12.2407
6.965 | -34.2719 | -12.2508
6.9711 | -34.2723 | -12.261
6.9772 | -34.2727 | -12.2708
6.9833 | -34.273 | -12.2812
6.9894 | -34.2734 | -12.2917
6.9955 | -34.2738 | -12.3018
7.0016 | -34.2741 | -12.3123
7.0077 | -34.2745 | -12.3227
7.0138 | -34.2749 | -12.3331
7.0199 | -34.2752 | -12.343
7.026 | -34.2755 | -12.353
7.0321 | -34.2758 | -12.3631
7.0382 | -34.2762 | -12.373
7.0443 | -34.2766 | -12.3834
7.0504 | -34.2769 | -12.3938
7.0565 | -34.2773 | -12.4042
7.0626 | -34.2777 | -12.4141
7.0687 | -34.278 | -12.4238
7.0748 | -34.2784 | -12.4337
7.0809 | -34.2787 | -12.4439
7.087 | -34.2791 | -12.4541
7.0932 | -34.2794 | -12.4641
7.0993 | -34.2798 | -12.4743
7.1054 | -34.2801 | -12.4841
7.1115 | -34.2804 | -12.4941
7.1176 | -34.2807 | -12.504
7.1237 | -34.281 | -12.514
7.1298 | -34.2813 | -12.5236
7.1359 | -34.2816 | -12.5334
7.142 | -34.2819 | -12.5435
7.1481 | -34.2823 | -12.5532
7.1542 | -34.2826 | -12.5631
7.1603 | -34.2829 | -12.573
7.1664 | -34.2833 | -12.5831
7.1725 | -34.2837 | -12.5929
7.1786 | -34.2841 | -12.603
7.1847 | -34.2844 | -12.6129
7.1908 | -34.2847 | -12.623
7.1969 | -34.2851 | -12.6329
7.203 | -34.2854 | -12.643
7.2091 | -34.2857 | -12.6527
7.2152 | -34.2861 | -12.6627
7.2213 | -34.2864 | -12.6724
7.2274 | -34.2867 | -12.6824
7.2335 | -34.2871 | -12.6923
7.2397 | -34.2874 | -12.7022
7.2458 | -34.2877 | -12.7117
7.2519 | -34.2881 | -12.7216
7.258 | -34.2884 | -12.7318
7.2641 | -34.2888 | -12.7416
7.2702 | -34.2891 | -12.7519
7.2763 | -34.2895 | -12.7616
7.2824 | -34.2899 | -12.7714
7.2885 | -34.2902 | -12.7813
7.2946 | -34.2906 | -12.7914
7.3007 | -34.2908 | -12.8013
7.3068 | -34.2912 | -12.8115
7.3129 | -34.2915 | -12.8212
7.319 | -34.2919 | -12.8312
7.3251 | -34.2922 | -12.8408
7.3312 | -34.2926 | -12.8503
7.3373 | -34.293 | -12.8599
7.3434 | -34.2934 | -12.8695
7.3495 | -34.2937 | -12.8789
7.3556 | -34.294 | -12.8892
7.3617 | -34.2943 | -12.8995
7.3678 | -34.2947 | -12.9095
7.3739 | -34.2951 | -12.919
7.3801 | -34.2954 | -12.9287
7.3862 | -34.2957 | -12.9385
7.3923 | -34.296 | -12.9483
7.3984 | -34.2964 | -12.958
7.4045 | -34.2967 | -12.9677
7.4106 | -34.2971 | -12.9774
7.4167 | -34.2974 | -12.9875
7.4228 | -34.2978 | -12.9972
7.4289 | -34.2981 | -13.0068
7.435 | -34.2985 | -13.0163
7.4411 | -34.2988 | -13.0258
7.4472 | -34.2992 | -13.0357
7.4533 | -34.2996 | -13.0455
7.4594 | -34.2999 | -13.0556
7.4655 | -34.3003 | -13.0653
7.4716 | -34.3005 | -13.0754
7.4777 | -34.3009 | -13.0846
7.4838 | -34.3013 | -13.0946
7.4899 | -34.3017 | -13.1044
7.496 | -34.302 | -13.1142
7.5021 | -34.3024 | -13.1238
7.5082 | -34.3027 | -13.1337
7.5143 | -34.303 | -13.1433
7.5204 | -34.3034 | -13.1528
7.5266 | -34.3037 | -13.1626
7.5327 | -34.304 | -13.1722
7.5388 | -34.3043 | -13.1819
7.5449 | -34.3046 | -13.1919
7.551 | -34.305 | -13.2018
7.5571 | -34.3053 | -13.2115
7.5632 | -34.3056 | -13.2212
7.5693 | -34.306 | -13.231
7.5754 | -34.3064 | -13.2408
7.5815 | -34.3067 | -13.2502
7.5876 | -34.307 | -13.2599
7.5937 | -34.3073 | -13.2695
7.5998 | -34.3077 | -13.2794
7.6059 | -34.308 | -13.2891
7.612 | -34.3084 | -13.2987
7.6181 | -34.3087 | -13.3084
7.6242 | -34.309 | -13.3182
7.6303 | -34.3094 | -13.3277
7.6364 | -34.3097 | -13.3377
7.6425 | -34.3101 | -13.3477
7.6486 | -34.3104 | -13.3575
7.6547 | -34.3108 | -13.3673
7.6608 | -34.3112 | -13.3771
7.667 | -34.3115 | -13.3862
7.6731 | -34.3118 | -13.3961
7.6792 | -34.3121 | -13.406
7.6853 | -34.3124 | -13.4156
7.6914 | -34.3126 | -13.4255
7.6975 | -34.313 | -13.4354
7.7036 | -34.3133 | -13.4453
7.7097 | -34.3136 | -13.4549
7.7158 | -34.314 | -13.4645
7.7219 | -34.3143 | -13.4742
7.728 | -34.3146 | -13.484
7.7341 | -34.3149 | -13.4938
7.7402 | -34.3153 | -13.5034
7.7463 | -34.3156 | -13.5132
7.7524 | -34.3158 | -13.5233
7.7585 | -34.3162 | -13.5331
7.7646 | -34.3165 | -13.5426
7.7707 | -34.3167 | -13.5521
7.7768 | -34.317 | -13.5617
7.7829 | -34.3173 | -13.5715
7.789 | -34.3176 | -13.581
7.7951 | -34.3179 | -13.5911
7.8012 | -34.3182 | -13.6012
7.8073 | -34.3185 | -13.6108
7.8135 | -34.3188 | -13.6206
7.8196 | -34.3192 | -13.6303
7.8257 | -34.3195 | -13.6404
7.8318 | -34.3198 | -13.6505
7.8379 | -34.3202 | -13.6598
7.844 | -34.3204 | -13.6691
7.8501 | -34.3206 | -13.6791
7.8562 | -34.3209 | -13.6895
7.8623 | -34.3213 | -13.6993
7.8684 | -34.3216 | -13.7087
7.8745 | -34.3219 | -13.7183
7.8806 | -34.3221 | -13.7276
7.8867 | -34.3225 | -13.7372
7.8928 | -34.3228 | -13.7467
7.8989 | -34.3232 | -13.756
7.905 | -34.3235 | -13.7661
7.9111 | -34.3238 | -13.7759
7.9172 | -34.3242 | -13.7853
7.9233 | -34.3244 | -13.7951
7.9294 | -34.3248 | -13.805
7.9355 | -34.3251 | -13.8142
7.9416 | -34.3253 | -13.8238
7.9477 | -34.3256 | -13.8331
7.9539 | -34.326 | -13.8424
7.96 | -34.3263 | -13.8524
7.9661 | -34.3266 | -13.8621
7.9722 | -34.3269 | -13.8719
7.9783 | -34.3272 | -13.8816
7.9844 | -34.3276 | -13.8917
7.9905 | -34.3279 | -13.9013
7.9966 | -34.3281 | -13.911
8.0027 | -34.3284 | -13.921
8.0088 | -34.3287 | -13.9303
8.0149 | -34.329 | -13.9397
8.021 | -34.3293 | -13.9494
8.0271 | -34.3297 | -13.9596
8.0332 | -34.3299 | -13.9692
8.0393 | -34.3303 | -13.9791
8.0454 | -34.3305 | -13.989
8.0515 | -34.3309 | -13.9987
8.0576 | -34.3311 | -14.0087
8.0637 | -34.3314 | -14.0178
8.0698 | -34.3317 | -14.0271
8.0759 | -34.332 | -14.0365
8.082 | -34.3324 | -14.0463
8.0881 | -34.3326 | -14.0559
8.0942 | -34.3329 | -14.066
8.1004 | -34.3332 | -14.0761
8.1065 | -34.3336 | -14.0857
8.1126 | -34.3339 | -14.0953
8.1187 | -34.3342 | -14.1051
8.1248 | -34.3345 | -14.115
8.1309 | -34.3348 | -14.1247
8.137 | -34.3351 | -14.1342
8.1431 | -34.3354 | -14.1441
8.1492 | -34.3357 | -14.1539
8.1553 | -34.3361 | -14.1636
8.1614 | -34.3363 | -14.1736
8.1675 | -34.3366 | -14.1832
8.1736 | -34.3369 | -14.1932
8.1797 | -34.3372 | -14.203
8.1858 | -34.3375 | -14.2127
8.1919 | -34.3377 | -14.2223
8.198 | -34.338 | -14.2316
8.2041 | -34.3384 | -14.2417
8.2102 | -34.3387 | -14.2514
8.2163 | -34.3391 | -14.2611
8.2224 | -34.3393 | -14.271
8.2285 | -34.3397 | -14.2805
8.2346 | -34.34 | -14.2901
8.2408 | -34.3403 | -14.2999
8.2469 | -34.3406 | -14.3095
8.253 | -34.3409 | -14.3192
8.2591 | -34.3412 | -14.3286
8.2652 | -34.3416 | -14.3381
8.2713 | -34.3419 | -14.3479
8.2774 | -34.3422 | -14.3577
8.2835 | -34.3425 | -14.3678
8.2896 | -34.3428 | -14.377
8.2957 | -34.3431 | -14.3867
8.3018 | -34.3434 | -14.3964
8.3079 | -34.3438 | -14.406
8.314 | -34.3441 | -14.4156
8.3201 | -34.3444 | -14.4252
8.3262 | -34.3447 | -14.4348
8.3323 | -34.345 | -14.4445
8.3384 | -34.3454 | -14.4546
8.3445 | -34.3457 | -14.4639
8.3506 | -34.346 | -14.4738
8.3567 | -34.3463 | -14.4838
8.3628 | -34.3466 | -14.4934
8.3689 | -34.3469 | -14.5031
8.375 | -34.3472 | -14.5128
8.3812 | -34.3475 | -14.5224
8.3873 | -34.3478 | -14.5322
8.3934 | -34.3481 | -14.5419
8.3995 | -34.3484 | -14.5517
8.4056 | -34.3487 | -14.5615
8.4117 | -34.349 | -14.5714
8.4178 | -34.3494 | -14.5807
8.4239 | -34.3497 | -14.5901
8.43 | -34.35 | -14.5998
8.4361 | -34.3503 | -14.6099
8.4422 | -34.3506 | -14.6196
8.4483 | -34.3509 | -14.6292
8.4544 | -34.3512 | -14.6392
8.4605 | -34.3515 | -14.6488
8.4666 | -34.3518 | -14.6586
8.4727 | -34.3521 | -14.6683
8.4788 | -34.3524 | -14.6778
8.4849 | -34.3527 | -14.6878
8.491 | -34.353 | -14.6973
8.4971 | -34.3533 | -14.7071
8.5032 | -34.3536 | -14.7168
8.5093 | -34.3539 | -14.7265
8.5154 | -34.3542 | -14.736
8.5215 | -34.3545 | -14.7457
8.5277 | -34.3548 | -14.7555
8.5338 | -34.3551 | -14.7649
8.5399 | -34.3554 | -14.7749
8.546 | -34.3558 | -14.7843
8.5521 | -34.3561 | -14.7941
8.5582 | -34.3564 | -14.8039
8.5643 | -34.3567 | -14.8134
8.5704 | -34.357 | -14.8228
8.5765 | -34.3573 | -14.8324
8.5826 | -34.3577 | -14.8421
8.5887 | -34.358 | -14.8517
8.5948 | -34.3582 | -14.861
8.6009 | -34.3586 | -14.8705
8.607 | -34.3589 | -14.8805
8.6131 | -34.3591 | -14.8904
8.6192 | -34.3595 | -14.8997
8.6253 | -34.3598 | -14.9095
8.6314 | -34.3601 | -14.9192
8.6375 | -34.3604 | -14.929
8.6436 | -34.3608 | -14.9387
8.6497 | -34.3611 | -14.9485
8.6558 | -34.3614 | -14.9579
8.6619 | -34.3617 | -14.9678
8.6681 | -34.362 | -14.9774
8.6742 | -34.3623 | -14.9873
8.6803 | -34.3627 | -14.9972
8.6864 | -34.363 | -15.0072
8.6925 | -34.3634 | -15.0169
8.6986 | -34.3637 | -15.0269
8.7047 | -34.364 | -15.0366
8.7108 | -34.3643 | -15.0464
8.7169 | -34.3646 | -15.0559
8.723 | -34.365 | -15.0656
8.7291 | -34.3653 | -15.0751
8.7352 | -34.3656 | -15.0851
8.7413 | -34.3659 | -15.0947
8.7474 | -34.3662 | -15.1041
8.7535 | -34.3666 | -15.1139
8.7596 | -34.367 | -15.1242
8.7657 | -34.3673 | -15.1338
8.7718 | -34.3676 | -15.1432
8.7779 | -34.3679 | -15.1535
8.784 | -34.3682 | -15.163
8.7901 | -34.3684 | -15.1732
8.7962 | -34.3688 | -15.183
8.8023 | -34.3692 | -15.1923
8.8084 | -34.3695 | -15.2022
8.8146 | -34.3698 | -15.2118
8.8207 | -34.3702 | -15.2215
8.8268 | -34.3705 | -15.2311
8.8329 | -34.3708 | -15.2409
8.839 | -34.3712 | -15.2507
8.8451 | -34.3715 | -15.2602
8.8512 | -34.3719 | -15.2699
8.8573 | -34.3722 | -15.2795
8.8634 | -34.3726 | -15.2891
8.8695 | -34.3729 | -15.2989
8.8756 | -34.3732 | -15.3084
8.8817 | -34.3736 | -15.3183
8.8878 | -34.3738 | -15.3283
8.8939 | -34.3741 | -15.3377
8.9 | -34.3744 | -15.3471
8.9061 | -34.3748 | -15.3568
8.9122 | -34.3751 | -15.3664
8.9183 | -34.3755 | -15.376
8.9244 | -34.3758 | -15.3854
8.9305 | -34.3762 | -15.3952
8.9366 | -34.3765 | -15.4047
8.9427 | -34.3769 | -15.4145
8.9488 | -34.3772 | -15.4242
8.955 | -34.3776 | -15.4338
8.9611 | -34.3779 | -15.4437
8.9672 | -34.3783 | -15.4538
8.9733 | -34.3786 | -15.4633
8.9794 | -34.379 | -15.4731
8.9855 | -34.3793 | -15.4827
8.9916 | -34.3796 | -15.4923
8.9977 | -34.38 | -15.5019
9.0038 | -34.3804 | -15.5118
9.0099 | -34.3807 | -15.5216
9.016 | -34.3811 | -15.5312
9.0221 | -34.3814 | -15.5412
9.0282 | -34.3818 | -15.551
9.0343 | -34.3822 | -15.5609
9.0404 | -34.3826 | -15.5705
9.0465 | -34.3829 | -15.5806
9.0526 | -34.3833 | -15.5902
9.0587 | -34.3837 | -15.5998
9.0648 | -34.3841 | -15.6094
9.0709 | -34.3844 | -15.6189
9.077 | -34.3847 | -15.6291
9.0831 | -34.3851 | -15.6389
9.0892 | -34.3854 | -15.6485
9.0953 | -34.3857 | -15.6584
9.1015 | -34.386 | -15.6681
9.1076 | -34.3864 | -15.6778
9.1137 | -34.3868 | -15.6875
9.1198 | -34.3872 | -15.6968
9.1259 | -34.3876 | -15.7065
9.132 | -34.388 | -15.7165
9.1381 | -34.3884 | -15.7261
9.1442 | -34.3888 | -15.7353
9.1503 | -34.3891 | -15.7452
9.1564 | -34.3895 | -15.7549
9.1625 | -34.3898 | -15.7646
9.1686 | -34.3902 | -15.7743
9.1747 | -34.3906 | -15.7842
9.1808 | -34.391 | -15.7941
9.1869 | -34.3913 | -15.8041
9.193 | -34.3916 | -15.8139
9.1991 | -34.392 | -15.8234
9.2052 | -34.3923 | -15.8329
9.2113 | -34.3928 | -15.8429
9.2174 | -34.3931 | -15.8533
9.2235 | -34.3935 | -15.8627
9.2296 | -34.3939 | -15.8725
9.2357 | -34.3943 | -15.8822
9.2419 | -34.3946 | -15.8919
9.248 | -34.3951 | -15.9018
9.2541 | -34.3954 | -15.9118
9.2602 | -34.3958 | -15.9216
9.2663 | -34.3962 | -15.9312
9.2724 | -34.3965 | -15.9409
9.2785 | -34.3969 | -15.9506
9.2846 | -34.3973 | -15.9605
9.2907 | -34.3976 | -15.9702
9.2968 | -34.398 | -15.98
9.3029 | -34.3983 | -15.9898
9.309 | -34.3987 | -15.9995
9.3151 | -34.3991 | -16.0091
9.3212 | -34.3995 | -16.0189
9.3273 | -34.3999 | -16.0285
9.3334 | -34.4003 | -16.0381
9.3395 | -34.4008 | -16.0477
9.3456 | -34.4012 | -16.0572
9.3517 | -34.4016 | -16.0666
9.3578 | -34.402 | -16.0761
9.3639 | -34.4024 | -16.086
9.37 | -34.4027 | -16.0957
9.3761 | -34.4031 | -16.1054
9.3822 | -34.4035 | -16.115
9.3884 | -34.4039 | -16.1247
9.3945 | -34.4043 | -16.1344
9.4006 | -34.4047 | -16.1442
9.4067 | -34.405 | -16.154
9.4128 | -34.4055 | -16.1636
9.4189 | -34.4059 | -16.1735
9.425 | -34.4063 | -16.183
9.4311 | -34.4067 | -16.1927
9.4372 | -34.407 | -16.2024
9.4433 | -34.4074 | -16.2122
9.4494 | -34.4078 | -16.2219
9.4555 | -34.4082 | -16.2315
9.4616 | -34.4086 | -16.2412
9.4677 | -34.4091 | -16.2508
9.4738 | -34.4094 | -16.2609
9.4799 | -34.4098 | -16.2705
9.486 | -34.4101 | -16.2805
9.4921 | -34.4105 | -16.2906
9.4982 | -34.4109 | -16.2998
9.5043 | -34.4113 | -16.3092
9.5104 | -34.4117 | -16.3188
9.5165 | -34.4121 | -16.3284
9.5226 | -34.4125 | -16.3383
9.5288 | -34.413 | -16.348
9.5349 | -34.4134 | -16.3576
9.541 | -34.4138 | -16.3673
9.5471 | -34.4141 | -16.3771
9.5532 | -34.4146 | -16.3868
9.5593 | -34.415 | -16.3965
9.5654 | -34.4154 | -16.4064
9.5715 | -34.4158 | -16.4163
9.5776 | -34.4162 | -16.4262
9.5837 | -34.4166 | -16.4356
9.5898 | -34.4169 | -16.4455
9.5959 | -34.4173 | -16.4551
9.602 | -34.4177 | -16.4648
9.6081 | -34.4181 | -16.4747
9.6142 | -34.4185 | -16.4843
9.6203 | -34.4188 | -16.4941
9.6264 | -34.4192 | -16.5034
9.6325 | -34.4195 | -16.5134
9.6386 | -34.42 | -16.523
9.6447 | -34.4203 | -16.5326
9.6508 | -34.4208 | -16.5424
9.6569 | -34.4211 | -16.552
9.663 | -34.4216 | -16.5618
9.6691 | -34.422 | -16.5715
9.6753 | -34.4225 | -16.5811
9.6814 | -34.4229 | -16.5905
9.6875 | -34.4234 | -16.6
9.6936 | -34.4238 | -16.6095
9.6997 | -34.4242 | -16.6194
9.7058 | -34.4246 | -16.6291
9.7119 | -34.425 | -16.6387
9.718 | -34.4254 | -16.6488
9.7241 | -34.4257 | -16.6585
9.7302 | -34.4261 | -16.6683
9.7363 | -34.4266 | -16.6778
9.7424 | -34.427 | -16.6873
9.7485 | -34.4274 | -16.6969
9.7546 | -34.4278 | -16.7063
9.7607 | -34.4282 | -16.7161
9.7668 | -34.4286 | -16.7262
9.7729 | -34.429 | -16.7356
9.779 | -34.4295 | -16.7453
9.7851 | -34.4298 | -16.7552
9.7912 | -34.4303 | -16.7653
9.7973 | -34.4306 | -16.7751
9.8034 | -34.4311 | -16.785
9.8095 | -34.4315 | -16.7946
9.8157 | -34.4319 | -16.8043
9.8218 | -34.4323 | -16.8136
9.8279 | -34.4327 | -16.8234
9.834 | -34.4331 | -16.8329
9.8401 | -34.4335 | -16.8425
9.8462 | -34.434 | -16.8519
9.8523 | -34.4344 | -16.8615
9.8584 | -34.4348 | -16.8714
9.8645 | -34.4352 | -16.881
9.8706 | -34.4356 | -16.8906
9.8767 | -34.4359 | -16.8998
9.8828 | -34.4364 | -16.9091
9.8889 | -34.4368 | -16.919
9.895 | -34.4372 | -16.9283
9.9011 | -34.4376 | -16.9374
9.9072 | -34.438 | -16.9467
9.9133 | -34.4384 | -16.9562
9.9194 | -34.4389 | -16.9656
9.9255 | -34.4393 | -16.9753
9.9316 | -34.4397 | -16.9847
9.9377 | -34.4402 | -16.9941
9.9438 | -34.4406 | -17.0036
9.9499 | -34.441 | -17.0133
9.956 | -34.4414 | -17.0227
9.9622 | -34.4418 | -17.0322
9.9683 | -34.4422 | -17.0419
9.9744 | -34.4426 | -17.0514
9.9805 | -34.4431 | -17.061
9.9866 | -34.4435 | -17.0702
9.9927 | -34.4439 | -17.08
9.9988 | -34.4443 | -17.0894
10.0049 | -34.4447 | -17.0989
10.011 | -34.4452 | -17.1084
10.0171 | -34.4456 | -17.1179
10.0232 | -34.446 | -17.1272
10.0293 | -34.4465 | -17.1366
10.0354 | -34.447 | -17.1459
10.0415 | -34.4474 | -17.1554
10.0476 | -34.4478 | -17.1652
10.0537 | -34.4482 | -17.175
10.0598 | -34.4486 | -17.1847
10.0659 | -34.449 | -17.1942
10.072 | -34.4495 | -17.2036
10.0781 | -34.4499 | -17.2131
10.0842 | -34.4503 | -17.2228
10.0903 | -34.4507 | -17.2325
10.0964 | -34.4512 | -17.2417
10.1026 | -34.4517 | -17.2512
10.1087 | -34.4521 | -17.2608
10.1148 | -34.4525 | -17.2704
10.1209 | -34.453 | -17.2797
10.127 | -34.4534 | -17.289
10.1331 | -34.4538 | -17.2983
10.1392 | -34.4543 | -17.3077
10.1453 | -34.4547 | -17.3172
10.1514 | -34.4551 | -17.3264
10.1575 | -34.4555 | -17.3358
10.1636 | -34.456 | -17.3445
10.1697 | -34.4564 | -17.354
10.1758 | -34.4569 | -17.3633
10.1819 | -34.4573 | -17.3728
10.188 | -34.4577 | -17.3821
10.1941 | -34.4581 | -17.3916
10.2002 | -34.4585 | -17.4005
10.2063 | -34.4589 | -17.4101
10.2124 | -34.4593 | -17.4196
10.2185 | -34.4597 | -17.429
10.2246 | -34.4602 | -17.4384
10.2307 | -34.4606 | -17.448
10.2368 | -34.461 | -17.4565
10.2429 | -34.4615 | -17.4656
10.2491 | -34.4619 | -17.4748
10.2552 | -34.4623 | -17.4846
10.2613 | -34.4628 | -17.4943
10.2674 | -34.4631 | -17.5039
10.2735 | -34.4635 | -17.5136
10.2796 | -34.464 | -17.5231
10.2857 | -34.4644 | -17.5326
10.2918 | -34.4648 | -17.5421
10.2979 | -34.4652 | -17.5517
10.304 | -34.4657 | -17.5613
10.3101 | -34.4662 | -17.5707
10.3162 | -34.4666 | -17.5802
10.3223 | -34.467 | -17.5896
10.3284 | -34.4675 | -17.5989
10.3345 | -34.4679 | -17.6081
10.3406 | -34.4683 | -17.6174
10.3467 | -34.4687 | -17.6268
10.3528 | -34.4691 | -17.6359
10.3589 | -34.4695 | -17.6455
10.365 | -34.4699 | -17.6552
10.3711 | -34.4703 | -17.6648
10.3772 | -34.4707 | -17.6743
10.3833 | -34.4711 | -17.6837
10.3895 | -34.4715 | -17.6929
10.3956 | -34.472 | -17.7021
10.4017 | -34.4724 | -17.7114
10.4078 | -34.4728 | -17.7208
10.4139 | -34.4733 | -17.7301
10.42 | -34.4737 | -17.7395
10.4261 | -34.4742 | -17.7488
10.4322 | -34.4746 | -17.758
10.4383 | -34.475 | -17.7673
10.4444 | -34.4754 | -17.7763
10.4505 | -34.4759 | -17.7857
10.4566 | -34.4762 | -17.795
10.4627 | -34.4767 | -17.8043
10.4688 | -34.4771 | -17.8131
10.4749 | -34.4776 | -17.8223
10.481 | -34.478 | -17.8317
10.4871 | -34.4784 | -17.8411
10.4932 | -34.4789 | -17.8504
10.4993 | -34.4794 | -17.8596
10.5054 | -34.4798 | -17.869
10.5115 | -34.4802 | -17.8783
10.5176 | -34.4807 | -17.8874
10.5237 | -34.4811 | -17.8966
10.5298 | -34.4815 | -17.9056
10.536 | -34.482 | -17.9152
10.5421 | -34.4824 | -17.924
10.5482 | -34.4828 | -17.9336
10.5543 | -34.4833 | -17.9428
10.5604 | -34.4837 | -17.9517
10.5665 | -34.4842 | -17.9608
10.5726 | -34.4846 | -17.9699
10.5787 | -34.485 | -17.9793
10.5848 | -34.4855 | -17.9882
10.5909 | -34.4859 | -17.9975
10.597 | -34.4863 | -18.0068
10.6031 | -34.4867 | -18.0162
10.6092 | -34.4872 | -18.0257
10.6153 | -34.4876 | -18.0352
10.6214 | -34.4881 | -18.0444
10.6275 | -34.4885 | -18.0535
10.6336 | -34.4889 | -18.0631
10.6397 | -34.4893 | -18.0723
10.6458 | -34.4897 | -18.0813
10.6519 | -34.4901 | -18.0905
10.658 | -34.4905 | -18.0995
10.6641 | -34.491 | -18.1088
10.6702 | -34.4914 | -18.1183
10.6764 | -34.4919 | -18.1274
10.6825 | -34.4924 | -18.1368
10.6886 | -34.4928 | -18.1462
10.6947 | -34.4932 | -18.1552
10.7008 | -34.4937 | -18.1643
10.7069 | -34.4941 | -18.1737
10.713 | -34.4946 | -18.1828
10.7191 | -34.495 | -18.1922
10.7252 | -34.4955 | -18.2011
10.7313 | -34.4959 | -18.2103
10.7374 | -34.4964 | -18.2193
10.7435 | -34.4968 | -18.2288
10.7496 | -34.4972 | -18.2378
10.7557 | -34.4976 | -18.247
10.7618 | -34.4981 | -18.2562
10.7679 | -34.4985 | -18.2653
10.774 | -34.4989 | -18.2745
10.7801 | -34.4994 | -18.2841
10.7862 | -34.4998 | -18.2932
10.7923 | -34.5003 | -18.3025
10.7984 | -34.5007 | -18.3115
10.8045 | -34.5011 | -18.3209
10.8106 | -34.5016 | -18.3301
10.8168 | -34.502 | -18.3393
10.8229 | -34.5024 | -18.3486
10.829 | -34.5029 | -18.3579
10.8351 | -34.5033 | -18.367
10.8412 | -34.5037 | -18.3761
10.8473 | -34.5041 | -18.385
10.8534 | -34.5046 | -18.3944
10.8595 | -34.505 | -18.4034
10.8656 | -34.5055 | -18.4124
10.8717 | -34.5059 | -18.4212
10.8778 | -34.5064 | -18.4302
10.8839 | -34.5068 | -18.4393
10.89 | -34.5072 | -18.4482
10.8961 | -34.5077 | -18.4575
10.9022 | -34.5081 | -18.4665
10.9083 | -34.5086 | -18.4758
10.9144 | -34.509 | -18.4848
10.9205 | -34.5095 | -18.494
10.9266 | -34.51 | -18.5034
10.9327 | -34.5105 | -18.5126
10.9388 | -34.5109 | -18.5217
10.9449 | -34.5113 | -18.5308
10.951 | -34.5117 | -18.5399
10.9571 | -34.5121 | -18.5491
10.9633 | -34.5126 | -18.5586
10.9694 | -34.513 | -18.5681
10.9755 | -34.5135 | -18.5773
10.9816 | -34.5139 | -18.5865
10.9877 | -34.5144 | -18.5956
10.9938 | -34.5148 | -18.605
10.9999 | -34.5153 | -18.614
11.006 | -34.5158 | -18.623
11.0121 | -34.5162 | -18.6325
11.0182 | -34.5166 | -18.6413
11.0243 | -34.5171 | -18.6505
11.0304 | -34.5175 | -18.6595
11.0365 | -34.5179 | -18.6685
11.0426 | -34.5184 | -18.6775
11.0487 | -34.5188 | -18.6867
11.0548 | -34.5193 | -18.6958
11.0609 | -34.5197 | -18.7052
11.067 | -34.5201 | -18.7145
11.0731 | -34.5206 | -18.7236
11.0792 | -34.521 | -18.7327
11.0853 | -34.5215 | -18.7415
11.0914 | -34.5219 | -18.7506
11.0975 | -34.5224 | -18.76
11.1037 | -34.5228 | -18.7688
11.1098 | -34.5232 | -18.7781
11.1159 | -34.5237 | -18.7873
11.122 | -34.5242 | -18.7967
11.1281 | -34.5246 | -18.8056
11.1342 | -34.525 | -18.8144
11.1403 | -34.5254 | -18.8238
11.1464 | -34.5259 | -18.8329
11.1525 | -34.5263 | -18.8421
11.1586 | -34.5268 | -18.8518
11.1647 | -34.5272 | -18.8606
11.1708 | -34.5277 | -18.8696
11.1769 | -34.528 | -18.8782
11.183 | -34.5285 | -18.8872
11.1891 | -34.529 | -18.8966
11.1952 | -34.5294 | -18.906
11.2013 | -34.5298 | -18.9151
11.2074 | -34.5303 | -18.9239
11.2135 | -34.5306 | -18.933
11.2196 | -34.531 | -18.9421
11.2257 | -34.5314 | -18.9511
11.2318 | -34.5318 | -18.9604
11.2379 | -34.5322 | -18.9694
11.244 | -34.5326 | -18.9785
11.2502 | -34.5331 | -18.9875
11.2563 | -34.5336 | -18.9966
11.2624 | -34.5339 | -19.0057
11.2685 | -34.5344 | -19.0146
11.2746 | -34.5348 | -19.0234
11.2807 | -34.5352 | -19.033
11.2868 | -34.5356 | -19.0424
11.2929 | -34.5361 | -19.0511
11.299 | -34.5365 | -19.0604
11.3051 | -34.5369 | -19.0695
11.3112 | -34.5373 | -19.0787
11.3173 | -34.5378 | -19.0878
11.3234 | -34.5382 | -19.097
11.3295 | -34.5386 | -19.1059
11.3356 | -34.5391 | -19.1148
11.3417 | -34.5395 | -19.1241
11.3478 | -34.5399 | -19.1331
11.3539 | -34.5403 | -19.1421
11.36 | -34.5407 | -19.151
11.3661 | -34.5412 | -19.16
11.3722 | -34.5416 | -19.1689
11.3783 | -34.5421 | -19.1779
11.3844 | -34.5424 | -19.187
11.3906 | -34.5429 | -19.1963
11.3967 | -34.5433 | -19.2052
11.4028 | -34.5437 | -19.2145
11.4089 | -34.5441 | -19.2234
11.415 | -34.5445 | -19.2323
11.4211 | -34.5449 | -19.2411
11.4272 | -34.5453 | -19.2503
11.4333 | -34.5458 | -19.2594
11.4394 | -34.5462 | -19.2687
11.4455 | -34.5466 | -19.2778
11.4516 | -34.547 | -19.2867
11.4577 | -34.5474 | -19.296
11.4638 | -34.5479 | -19.3046
11.4699 | -34.5483 | -19.314
11.476 | -34.5487 | -19.3235
11.4821 | -34.5492 | -19.3323
11.4882 | -34.5496 | -19.3417
11.4943 | -34.55 | -19.3506
11.5004 | -34.5505 | -19.3596
11.5065 | -34.551 | -19.3687
11.5126 | -34.5514 | -19.3774
11.5187 | -34.5518 | -19.3864
11.5248 | -34.5523 | -19.3955
11.5309 | -34.5527 | -19.4045
11.5371 | -34.5531 | -19.4134
11.5432 | -34.5534 | -19.4227
11.5493 | -34.5539 | -19.4318
11.5554 | -34.5543 | -19.4407
11.5615 | -34.5547 | -19.4499
11.5676 | -34.5552 | -19.4592
11.5737 | -34.5556 | -19.4679
11.5798 | -34.556 | -19.4767
11.5859 | -34.5565 | -19.4856
11.592 | -34.5569 | -19.4946
11.5981 | -34.5574 | -19.5035
11.6042 | -34.5578 | -19.5126
11.6103 | -34.5582 | -19.5215
11.6164 | -34.5586 | -19.5303
11.6225 | -34.559 | -19.5393
11.6286 | -34.5595 | -19.5484
11.6347 | -34.5599 | -19.5577
11.6408 | -34.5603 | -19.5667
11.6469 | -34.5608 | -19.5759
11.653 | -34.5612 | -19.5847
11.6591 | -34.5617 | -19.5937
11.6652 | -34.5621 | -19.6028
11.6713 | -34.5625 | -19.6118
11.6775 | -34.563 | -19.6205
11.6836 | -34.5634 | -19.6296
11.6897 | -34.5639 | -19.6384
11.6958 | -34.5643 | -19.6478
11.7019 | -34.5647 | -19.6569
11.708 | -34.5651 | -19.6661
11.7141 | -34.5656 | -19.6752
11.7202 | -34.566 | -19.6839
11.7263 | -34.5664 | -19.693
11.7324 | -34.5669 | -19.7021
11.7385 | -34.5673 | -19.7113
11.7446 | -34.5677 | -19.7204
11.7507 | -34.5682 | -19.7293
11.7568 | -34.5686 | -19.7384
11.7629 | -34.569 | -19.7475
11.769 | -34.5695 | -19.7564
11.7751 | -34.5699 | -19.7652
11.7812 | -34.5703 | -19.774
11.7873 | -34.5707 | -19.7833
11.7934 | -34.5712 | -19.7922
11.7995 | -34.5716 | -19.8009
11.8056 | -34.5721 | -19.8102
11.8117 | -34.5724 | -19.8193
11.8178 | -34.5729 | -19.8282
11.824 | -34.5734 | -19.837
11.8301 | -34.5738 | -19.8455
11.8362 | -34.5742 | -19.8541
11.8423 | -34.5746 | -19.8632
11.8484 | -34.5751 | -19.8721
11.8545 | -34.5755 | -19.8814
11.8606 | -34.576 | -19.8903
11.8667 | -34.5764 | -19.8994
11.8728 | -34.5768 | -19.9083
11.8789 | -34.5772 | -19.9177
11.885 | -34.5776 | -19.9266
11.8911 | -34.5781 | -19.9357
11.8972 | -34.5785 | -19.9443
11.9033 | -34.5789 | -19.9528
11.9094 | -34.5793 | -19.9614
11.9155 | -34.5797 | -19.9704
11.9216 | -34.5802 | -19.9793
11.9277 | -34.5806 | -19.9879
11.9338 | -34.581 | -19.997
11.9399 | -34.5814 | -20.0063
11.946 | -34.5819 | -20.0152
11.9521 | -34.5823 | -20.0241
11.9582 | -34.5828 | -20.0328
11.9644 | -34.5832 | -20.0416
11.9705 | -34.5837 | -20.0507
11.9766 | -34.5841 | -20.0596
11.9827 | -34.5845 | -20.0684
11.9888 | -34.5849 | -20.0771
11.9949 | -34.5853 | -20.086
12.001 | -34.5858 | -20.095
12.0071 | -34.5863 | -20.1036
12.0132 | -34.5867 | -20.1122
12.0193 | -34.5871 | -20.1214
12.0254 | -34.5875 | -20.1301
12.0315 | -34.588 | -20.1389
12.0376 | -34.5884 | -20.1475
12.0437 | -34.5888 | -20.1562
12.0498 | -34.5892 | -20.1653
12.0559 | -34.5897 | -20.1742
12.062 | -34.5901 | -20.1831
12.0681 | -34.5905 | -20.1919
12.0742 | -34.591 | -20.2009
12.0803 | -34.5914 | -20.2097
12.0864 | -34.5919 | -20.2186
12.0925 | -34.5923 | -20.2275
12.0986 | -34.5928 | -20.2361
12.1047 | -34.5932 | -20.2449
12.1109 | -34.5937 | -20.2539
12.117 | -34.5941 | -20.2627
12.1231 | -34.5946 | -20.2715
12.1292 | -34.595 | -20.2805
12.1353 | -34.5955 | -20.2891
12.1414 | -34.5959 | -20.2981
12.1475 | -34.5963 | -20.307
12.1536 | -34.5967 | -20.316
12.1597 | -34.5972 | -20.3249
12.1658 | -34.5976 | -20.3337
12.1719 | -34.5981 | -20.3426
12.178 | -34.5986 | -20.3516
12.1841 | -34.599 | -20.3606
12.1902 | -34.5995 | -20.3695
12.1963 | -34.5999 | -20.3783
12.2024 | -34.6004 | -20.387
12.2085 | -34.6008 | -20.3959
12.2146 | -34.6012 | -20.4046
12.2207 | -34.6017 | -20.4135
12.2268 | -34.6021 | -20.4223
12.2329 | -34.6025 | -20.431
12.239 | -34.6029 | -20.4396
12.2451 | -34.6034 | -20.4481
12.2513 | -34.6039 | -20.4571
12.2574 | -34.6043 | -20.4662
12.2635 | -34.6047 | -20.4746
12.2696 | -34.6051 | -20.4832
12.2757 | -34.6056 | -20.4916
12.2818 | -34.606 | -20.5002
12.2879 | -34.6064 | -20.5089
12.294 | -34.6069 | -20.5174
12.3001 | -34.6073 | -20.5264
12.3062 | -34.6078 | -20.5352
12.3123 | -34.6082 | -20.544
12.3184 | -34.6087 | -20.5526
12.3245 | -34.6091 | -20.5615
12.3306 | -34.6095 | -20.57
12.3367 | -34.61 | -20.5786
12.3428 | -34.6105 | -20.5871
12.3489 | -34.6109 | -20.5953
12.355 | -34.6113 | -20.6041
12.3611 | -34.6118 | -20.6128
12.3672 | -34.6122 | -20.6214
12.3733 | -34.6126 | -20.6305
12.3794 | -34.613 | -20.6387
12.3855 | -34.6135 | -20.6472
12.3916 | -34.6139 | -20.6562
12.3978 | -34.6144 | -20.6652
12.4039 | -34.6147 | -20.6739
12.41 | -34.6152 | -20.6828
12.4161 | -34.6157 | -20.6913
12.4222 | -34.6161 | -20.6998
12.4283 | -34.6166 | -20.7085
12.4344 | -34.617 | -20.7173
12.4405 | -34.6175 | -20.726
12.4466 | -34.6179 | -20.7346
12.4527 | -34.6183 | -20.7434
12.4588 | -34.6188 | -20.7521
12.4649 | -34.6192 | -20.761
12.471 | -34.6197 | -20.7702
12.4771 | -34.6201 | -20.7783
12.4832 | -34.6206 | -20.7869
12.4893 | -34.621 | -20.7954
12.4954 | -34.6215 | -20.8039
12.5015 | -34.622 | -20.8123
12.5076 | -34.6224 | -20.8211
12.5137 | -34.6228 | -20.8302
12.5198 | -34.6232 | -20.8386
12.5259 | -34.6237 | -20.8471
12.532 | -34.6242 | -20.856
12.5382 | -34.6246 | -20.8645
12.5443 | -34.625 | -20.8732
12.5504 | -34.6255 | -20.8819
12.5565 | -34.626 | -20.8904
12.5626 | -34.6264 | -20.8993
12.5687 | -34.6269 | -20.9081
12.5748 | -34.6273 | -20.9167
12.5809 | -34.6277 | -20.9249
12.587 | -34.6281 | -20.9339
12.5931 | -34.6286 | -20.9426
12.5992 | -34.629 | -20.9514
12.6053 | -34.6295 | -20.9599
12.6114 | -34.63 | -20.9686
12.6175 | -34.6304 | -20.9774
12.6236 | -34.6308 | -20.9857
12.6297 | -34.6313 | -20.9942
12.6358 | -34.6318 | -21.0026
12.6419 | -34.6323 | -21.0114
12.648 | -34.6327 | -21.0203
12.6541 | -34.6332 | -21.0289
12.6602 | -34.6337 | -21.037
12.6663 | -34.634 | -21.0455
12.6724 | -34.6345 | -21.0541
12.6785 | -34.635 | -21.063
12.6847 | -34.6354 | -21.0717
12.6908 | -34.6359 | -21.0802
12.6969 | -34.6364 | -21.0888
12.703 | -34.6368 | -21.0975
12.7091 | -34.6372 | -21.1062
12.7152 | -34.6377 | -21.1147
12.7213 | -34.6381 | -21.1234
12.7274 | -34.6386 | -21.1321
12.7335 | -34.639 | -21.1401
12.7396 | -34.6394 | -21.1482
12.7457 | -34.6399 | -21.1569
12.7518 | -34.6403 | -21.1655
12.7579 | -34.6408 | -21.1744
12.764 | -34.6412 | -21.183
12.7701 | -34.6416 | -21.1913
12.7762 | -34.642 | -21.1999
12.7823 | -34.6425 | -21.2085
12.7884 | -34.6429 | -21.217
12.7945 | -34.6433 | -21.2253
12.8006 | -34.6438 | -21.2332
12.8067 | -34.6442 | -21.242
12.8128 | -34.6446 | -21.2501
12.8189 | -34.645 | -21.2584
12.8251 | -34.6454 | -21.2667
12.8312 | -34.6458 | -21.2753
12.8373 | -34.6463 | -21.2839
12.8434 | -34.6468 | -21.2924
12.8495 | -34.6472 | -21.3008
12.8556 | -34.6476 | -21.3092
12.8617 | -34.648 | -21.3174
12.8678 | -34.6484 | -21.326
12.8739 | -34.6489 | -21.3341
12.88 | -34.6493 | -21.3426
12.8861 | -34.6497 | -21.3513
12.8922 | -34.6501 | -21.3597
12.8983 | -34.6506 | -21.3679
12.9044 | -34.6511 | -21.3763
12.9105 | -34.6516 | -21.3846
12.9166 | -34.652 | -21.3934
12.9227 | -34.6524 | -21.4024
12.9288 | -34.6529 | -21.411
12.9349 | -34.6534 | -21.4198
12.941 | -34.6539 | -21.4284
12.9471 | -34.6544 | -21.4369
12.9532 | -34.6549 | -21.4454
12.9593 | -34.6553 | -21.4541
12.9654 | -34.6557 | -21.4627
12.9716 | -34.6562 | -21.4713
12.9777 | -34.6566 | -21.4794
12.9838 | -34.657 | -21.4877
12.9899 | -34.6574 | -21.4962
12.996 | -34.6578 | -21.5049
13.0021 | -34.6583 | -21.5134
13.0082 | -34.6587 | -21.5219
13.0143 | -34.6591 | -21.53
13.0204 | -34.6595 | -21.5381
13.0265 | -34.66 | -21.5463
13.0326 | -34.6604 | -21.5551
13.0387 | -34.6609 | -21.5637
13.0448 | -34.6613 | -21.5722
13.0509 | -34.6618 | -21.5807
13.057 | -34.6622 | -21.5887
13.0631 | -34.6628 | -21.597
13.0692 | -34.6632 | -21.6053
13.0753 | -34.6637 | -21.6137
13.0814 | -34.6641 | -21.6219
13.0875 | -34.6645 | -21.6299
13.0936 | -34.665 | -21.6382
13.0997 | -34.6654 | -21.6463
13.1058 | -34.6659 | -21.654
13.112 | -34.6663 | -21.6625
13.1181 | -34.6667 | -21.6709
13.1242 | -34.6671 | -21.6794
13.1303 | -34.6676 | -21.688
13.1364 | -34.668 | -21.6963
13.1425 | -34.6684 | -21.7046
13.1486 | -34.6689 | -21.7129
13.1547 | -34.6693 | -21.7209
13.1608 | -34.6698 | -21.7292
13.1669 | -34.6702 | -21.7374
13.173 | -34.6707 | -21.7457
13.1791 | -34.6711 | -21.7541
13.1852 | -34.6715 | -21.7628
13.1913 | -34.672 | -21.771
13.1974 | -34.6724 | -21.7794
13.2035 | -34.6728 | -21.7876
13.2096 | -34.6733 | -21.7959
13.2157 | -34.6737 | -21.8043
13.2218 | -34.6742 | -21.8128
13.2279 | -34.6747 | -21.8211
13.234 | -34.6751 | -21.8296
13.2401 | -34.6755 | -21.8378
13.2462 | -34.676 | -21.8461
13.2524 | -34.6764 | -21.8544
13.2585 | -34.6769 | -21.8627
13.2646 | -34.6773 | -21.8712
13.2707 | -34.6777 | -21.8797
13.2768 | -34.6782 | -21.8883
13.2829 | -34.6786 | -21.8967
13.289 | -34.679 | -21.9051
13.2951 | -34.6795 | -21.9132
13.3012 | -34.6799 | -21.9216
13.3073 | -34.6803 | -21.9303
13.3134 | -34.6807 | -21.9385
13.3195 | -34.6812 | -21.9467
13.3256 | -34.6817 | -21.9552
13.3317 | -34.6821 | -21.9633
13.3378 | -34.6825 | -21.9717
13.3439 | -34.6829 | -21.9801
13.35 | -34.6834 | -21.9884
13.3561 | -34.6838 | -21.9968
13.3622 | -34.6842 | -22.0051
13.3683 | -34.6847 | -22.0132
13.3744 | -34.6851 | -22.0215
13.3805 | -34.6855 | -22.03
13.3866 | -34.686 | -22.0382
13.3927 | -34.6864 | -22.0464
13.3989 | -34.6868 | -22.0545
13.405 | -34.6873 | -22.0627
13.4111 | -34.6878 | -22.071
13.4172 | -34.6882 | -22.079
13.4233 | -34.6886 | -22.0876
13.4294 | -34.689 | -22.0956
13.4355 | -34.6895 | -22.1044
13.4416 | -34.6899 | -22.1128
13.4477 | -34.6903 | -22.1211
13.4538 | -34.6907 | -22.1292
13.4599 | -34.6912 | -22.1377
13.466 | -34.6916 | -22.1465
13.4721 | -34.6921 | -22.1549
13.4782 | -34.6925 | -22.1632
13.4843 | -34.693 | -22.1715
13.4904 | -34.6935 | -22.1793
13.4965 | -34.6939 | -22.1881
13.5026 | -34.6943 | -22.1963
13.5087 | -34.6948 | -22.2044
13.5148 | -34.6952 | -22.2128
13.5209 | -34.6956 | -22.221
13.527 | -34.6961 | -22.2297
13.5331 | -34.6965 | -22.2378
13.5393 | -34.697 | -22.2462
13.5454 | -34.6974 | -22.2545
13.5515 | -34.6979 | -22.2627
13.5576 | -34.6983 | -22.2712
13.5637 | -34.6987 | -22.2794
13.5698 | -34.6992 | -22.2879
13.5759 | -34.6995 | -22.2958
13.582 | -34.6999 | -22.3042
13.5881 | -34.7004 | -22.3124
13.5942 | -34.7008 | -22.3208
13.6003 | -34.7013 | -22.3291
13.6064 | -34.7017 | -22.3376
13.6125 | -34.7022 | -22.3463
13.6186 | -34.7026 | -22.3545
13.6247 | -34.703 | -22.3627
13.6308 | -34.7035 | -22.3707
13.6369 | -34.7039 | -22.379
13.643 | -34.7044 | -22.3872
13.6491 | -34.7048 | -22.3954
13.6552 | -34.7052 | -22.4034
13.6613 | -34.7057 | -22.4119
13.6674 | -34.7061 | -22.4203
13.6735 | -34.7066 | -22.4287
13.6796 | -34.707 | -22.4368
13.6858 | -34.7074 | -22.4451
13.6919 | -34.7079 | -22.4534
13.698 | -34.7083 | -22.4618
13.7041 | -34.7087 | -22.4703
13.7102 | -34.7091 | -22.4787
13.7163 | -34.7096 | -22.4872
13.7224 | -34.7101 | -22.4954
13.7285 | -34.7105 | -22.5038
13.7346 | -34.7109 | -22.5121
13.7407 | -34.7113 | -22.5204
13.7468 | -34.7118 | -22.5288
13.7529 | -34.7122 | -22.5369
13.759 | -34.7126 | -22.5454
13.7651 | -34.713 | -22.5538
13.7712 | -34.7134 | -22.5623
13.7773 | -34.7139 | -22.5707
13.7834 | -34.7144 | -22.579
13.7895 | -34.7149 | -22.5872
13.7956 | -34.7153 | -22.5956
13.8017 | -34.7158 | -22.6038
13.8078 | -34.7162 | -22.6121
13.8139 | -34.7167 | -22.6201
13.82 | -34.7171 | -22.6281
13.8262 | -34.7176 | -22.6369
13.8323 | -34.718 | -22.6455
13.8384 | -34.7185 | -22.654
13.8445 | -34.7189 | -22.6624
13.8506 | -34.7194 | -22.6709
13.8567 | -34.7198 | -22.6795
13.8628 | -34.7202 | -22.688
13.8689 | -34.7206 | -22.6964
13.875 | -34.7211 | -22.7047
13.8811 | -34.7216 | -22.713
13.8872 | -34.722 | -22.7212
13.8933 | -34.7225 | -22.7295
13.8994 | -34.723 | -22.7376
13.9055 | -34.7234 | -22.7459
13.9116 | -34.7239 | -22.7539
13.9177 | -34.7243 | -22.762
13.9238 | -34.7248 | -22.7705
13.9299 | -34.7253 | -22.7788
13.936 | -34.7257 | -22.7869
13.9421 | -34.7261 | -22.7953
13.9482 | -34.7266 | -22.8035
13.9543 | -34.727 | -22.8116
13.9604 | -34.7275 | -22.8201
13.9665 | -34.728 | -22.8285
13.9727 | -34.7284 | -22.8367
13.9788 | -34.7288 | -22.8452
13.9849 | -34.7292 | -22.8536
13.991 | -34.7296 | -22.8623
13.9971 | -34.73 | -22.8709
14.0032 | -34.7304 | -22.8794
14.0093 | -34.7309 | -22.8876
14.0154 | -34.7313 | -22.8961
14.0215 | -34.7317 | -22.9041
14.0276 | -34.7322 | -22.9122
14.0337 | -34.7327 | -22.9208
14.0398 | -34.7331 | -22.9289
14.0459 | -34.7335 | -22.9374
14.052 | -34.7339 | -22.9457
14.0581 | -34.7343 | -22.9536
14.0642 | -34.7348 | -22.9621
14.0703 | -34.7352 | -22.9705
14.0764 | -34.7357 | -22.9791
14.0825 | -34.7361 | -22.9875
14.0886 | -34.7365 | -22.9953
14.0947 | -34.7369 | -23.0033
14.1008 | -34.7374 | -23.012
14.1069 | -34.7378 | -23.0203
14.1131 | -34.7382 | -23.0289
14.1192 | -34.7387 | -23.0369
14.1253 | -34.7391 | -23.0449
14.1314 | -34.7396 | -23.0531
14.1375 | -34.74 | -23.0616
14.1436 | -34.7405 | -23.0698
14.1497 | -34.741 | -23.0778
14.1558 | -34.7414 | -23.0862
14.1619 | -34.7418 | -23.0943
14.168 | -34.7423 | -23.1026
14.1741 | -34.7428 | -23.1106
14.1802 | -34.7432 | -23.1189
14.1863 | -34.7437 | -23.1273
14.1924 | -34.7441 | -23.1358
14.1985 | -34.7446 | -23.1438
14.2046 | -34.745 | -23.1519
14.2107 | -34.7455 | -23.1598
14.2168 | -34.746 | -23.1685
14.2229 | -34.7464 | -23.1767
14.229 | -34.7468 | -23.185
14.2351 | -34.7473 | -23.193
14.2412 | -34.7477 | -23.2012
14.2473 | -34.7483 | -23.2096
14.2534 | -34.7487 | -23.2178
14.2596 | -34.7491 | -23.226
14.2657 | -34.7496 | -23.2344
14.2718 | -34.75 | -23.2428
14.2779 | -34.7505 | -23.2508
14.284 | -34.7509 | -23.2591
14.2901 | -34.7514 | -23.2672
14.2962 | -34.7519 | -23.2752
14.3023 | -34.7523 | -23.2835
14.3084 | -34.7528 | -23.292
14.3145 | -34.7533 | -23.3003
14.3206 | -34.7537 | -23.3086
14.3267 | -34.7541 | -23.3168
14.3328 | -34.7546 | -23.325
14.3389 | -34.755 | -23.333
14.345 | -34.7554 | -23.3414
14.3511 | -34.7559 | -23.3494
14.3572 | -34.7563 | -23.3579
14.3633 | -34.7567 | -23.3661
14.3694 | -34.7571 | -23.3744
14.3755 | -34.7576 | -23.3823
14.3816 | -34.758 | -23.3905
14.3877 | -34.7585 | -23.3985
14.3938 | -34.7589 | -23.4071
14.4 | -34.7594 | -23.4151
14.4061 | -34.7598 | -23.4232
14.4122 | -34.7602 | -23.4317
14.4183 | -34.7606 | -23.4399
14.4244 | -34.761 | -23.4479
14.4305 | -34.7615 | -23.4565
14.4366 | -34.762 | -23.4648
14.4427 | -34.7624 | -23.4733
14.4488 | -34.7628 | -23.4812
14.4549 | -34.7633 | -23.4896
14.461 | -34.7638 | -23.4979
14.4671 | -34.7642 | -23.5062
14.4732 | -34.7647 | -23.514
14.4793 | -34.7651 | -23.5229
14.4854 | -34.7657 | -23.5308
14.4915 | -34.7661 | -23.5388
14.4976 | -34.7666 | -23.5471
14.5037 | -34.7671 | -23.5555
14.5098 | -34.7675 | -23.5635
14.5159 | -34.768 | -23.5719
14.522 | -34.7684 | -23.5802
14.5281 | -34.7689 | -23.5882
14.5342 | -34.7693 | -23.5967
14.5403 | -34.7698 | -23.6049
14.5465 | -34.7703 | -23.6135
14.5526 | -34.7707 | -23.6219
14.5587 | -34.7712 | -23.6302
14.5648 | -34.7716 | -23.6382
14.5709 | -34.7721 | -23.6466
14.577 | -34.7725 | -23.6549
14.5831 | -34.773 | -23.6632
14.5892 | -34.7735 | -23.671
14.5953 | -34.7739 | -23.6789
14.6014 | -34.7744 | -23.6871
14.6075 | -34.7749 | -23.695
14.6136 | -34.7753 | -23.703
14.6197 | -34.7757 | -23.7113
14.6258 | -34.7762 | -23.7196
14.6319 | -34.7767 | -23.7275
14.638 | -34.7772 | -23.7357
14.6441 | -34.7776 | -23.7436
14.6502 | -34.778 | -23.7515
14.6563 | -34.7785 | -23.7597
14.6624 | -34.779 | -23.7677
14.6685 | -34.7794 | -23.7756
14.6746 | -34.7799 | -23.7839
14.6807 | -34.7803 | -23.792
14.6869 | -34.7808 | -23.8001
14.693 | -34.7812 | -23.8081
14.6991 | -34.7816 | -23.8161
14.7052 | -34.7822 | -23.8242
14.7113 | -34.7826 | -23.8325
14.7174 | -34.7831 | -23.8405
14.7235 | -34.7835 | -23.8485
14.7296 | -34.784 | -23.8566
14.7357 | -34.7844 | -23.8645
14.7418 | -34.7849 | -23.8724
14.7479 | -34.7854 | -23.8807
14.754 | -34.7858 | -23.8887
14.7601 | -34.7863 | -23.8969
14.7662 | -34.7868 | -23.905
14.7723 | -34.7872 | -23.913
14.7784 | -34.7877 | -23.9213
14.7845 | -34.7882 | -23.9295
14.7906 | -34.7886 | -23.9378
14.7967 | -34.7891 | -23.9458
14.8028 | -34.7895 | -23.9538
14.8089 | -34.79 | -23.9617
14.815 | -34.7904 | -23.9697
14.8211 | -34.7908 | -23.9777
14.8272 | -34.7913 | -23.986
14.8334 | -34.7917 | -23.9939
14.8395 | -34.7922 | -24.002
14.8456 | -34.7927 | -24.01
14.8517 | -34.7931 | -24.0182
14.8578 | -34.7936 | -24.0262
14.8639 | -34.7941 | -24.0345
14.87 | -34.7946 | -24.0426
14.8761 | -34.795 | -24.0504
14.8822 | -34.7955 | -24.0587
14.8883 | -34.7959 | -24.0666
14.8944 | -34.7964 | -24.0747
14.9005 | -34.7969 | -24.0825
14.9066 | -34.7973 | -24.0905
14.9127 | -34.7978 | -24.0983
14.9188 | -34.7983 | -24.1067
14.9249 | -34.7987 | -24.1147
14.931 | -34.7992 | -24.1227
14.9371 | -34.7997 | -24.1307
14.9432 | -34.8002 | -24.1389
14.9493 | -34.8006 | -24.1469
14.9554 | -34.8011 | -24.1549
14.9615 | -34.8015 | -24.163
14.9676 | -34.802 | -24.1713
14.9738 | -34.8024 | -24.1792
14.9799 | -34.8029 | -24.1878
14.986 | -34.8034 | -24.1954
14.9921 | -34.8039 | -24.2032
14.9982 | -34.8043 | -24.2117
15.0043 | -34.8048 | -24.2196
15.0104 | -34.8052 | -24.2277
15.0165 | -34.8057 | -24.2359
15.0226 | -34.8062 | -24.2437
15.0287 | -34.8066 | -24.2522
15.0348 | -34.8071 | -24.2603
15.0409 | -34.8075 | -24.2687
15.047 | -34.808 | -24.277
15.0531 | -34.8086 | -24.2851
15.0592 | -34.8091 | -24.2933
15.0653 | -34.8096 | -24.3013
15.0714 | -34.81 | -24.309
15.0775 | -34.8105 | -24.3167
15.0836 | -34.811 | -24.3248
15.0897 | -34.8115 | -24.3328
15.0958 | -34.812 | -24.3407
15.1019 | -34.8124 | -24.3488
15.108 | -34.8129 | -24.3568
15.1141 | -34.8133 | -24.3648
15.1203 | -34.8138 | -24.3729
15.1264 | -34.8144 | -24.3811
15.1325 | -34.8149 | -24.389
15.1386 | -34.8153 | -24.3971
15.1447 | -34.8158 | -24.4053
15.1508 | -34.8163 | -24.4136
15.1569 | -34.8168 | -24.4219
15.163 | -34.8174 | -24.4294
15.1691 | -34.8179 | -24.4373
15.1752 | -34.8184 | -24.4453
15.1813 | -34.8188 | -24.4533
15.1874 | -34.8193 | -24.4615
15.1935 | -34.8199 | -24.4691
15.1996 | -34.8204 | -24.4774
15.2057 | -34.8208 | -24.4853
15.2118 | -34.8213 | -24.4933
15.2179 | -34.8218 | -24.5012
15.224 | -34.8222 | -24.5093
15.2301 | -34.8227 | -24.5171
15.2362 | -34.8232 | -24.525
15.2423 | -34.8237 | -24.5333
15.2484 | -34.8242 | -24.5413
15.2545 | -34.8247 | -24.5492
15.2607 | -34.8252 | -24.5572
15.2668 | -34.8257 | -24.5648
15.2729 | -34.8262 | -24.5727
15.279 | -34.8267 | -24.5808
15.2851 | -34.8272 | -24.5888
15.2912 | -34.8276 | -24.5967
15.2973 | -34.8281 | -24.6046
15.3034 | -34.8286 | -24.6124
15.3095 | -34.8291 | -24.6202
15.3156 | -34.8295 | -24.6282
15.3217 | -34.83 | -24.6358
15.3278 | -34.8305 | -24.6439
15.3339 | -34.8311 | -24.6516
15.34 | -34.8316 | -24.6593
15.3461 | -34.832 | -24.6672
15.3522 | -34.8325 | -24.6751
15.3583 | -34.833 | -24.683
15.3644 | -34.8335 | -24.6911
15.3705 | -34.834 | -24.6988
15.3766 | -34.8345 | -24.707
15.3827 | -34.835 | -24.7146
15.3888 | -34.8354 | -24.7223
15.3949 | -34.8359 | -24.7301
15.401 | -34.8364 | -24.7381
15.4072 | -34.8369 | -24.7457
15.4133 | -34.8374 | -24.7536
15.4194 | -34.8379 | -24.7616
15.4255 | -34.8384 | -24.7695
15.4316 | -34.8389 | -24.7774
15.4377 | -34.8394 | -24.7853
15.4438 | -34.8399 | -24.7929
15.4499 | -34.8404 | -24.8008
15.456 | -34.8408 | -24.8086
15.4621 | -34.8413 | -24.8166
15.4682 | -34.8418 | -24.8247
15.4743 | -34.8423 | -24.8325
15.4804 | -34.8428 | -24.8404
15.4865 | -34.8433 | -24.8482
15.4926 | -34.8438 | -24.8562
15.4987 | -34.8443 | -24.8643
15.5048 | -34.8448 | -24.8721
15.5109 | -34.8453 | -24.8796
15.517 | -34.8458 | -24.8876
15.5231 | -34.8463 | -24.8955
15.5292 | -34.8468 | -24.9032
15.5353 | -34.8473 | -24.911
15.5414 | -34.8478 | -24.9191
15.5476 | -34.8483 | -24.9271
15.5537 | -34.8488 | -24.9349
15.5598 | -34.8493 | -24.9426
15.5659 | -34.8498 | -24.9507
15.572 | -34.8503 | -24.9582
15.5781 | -34.8508 | -24.9661
15.5842 | -34.8514 | -24.9736
15.5903 | -34.8518 | -24.9811
15.5964 | -34.8524 | -24.9888
15.6025 | -34.8529 | -24.9965
15.6086 | -34.8534 | -25.0044
15.6147 | -34.8538 | -25.0123
15.6208 | -34.8544 | -25.0199
15.6269 | -34.8549 | -25.0276
15.633 | -34.8554 | -25.0355
15.6391 | -34.8559 | -25.0435
15.6452 | -34.8564 | -25.0512
15.6513 | -34.8569 | -25.0592
15.6574 | -34.8574 | -25.0668
15.6635 | -34.8579 | -25.0747
15.6696 | -34.8584 | -25.0823
15.6757 | -34.8589 | -25.0905
15.6818 | -34.8594 | -25.0986
15.688 | -34.86 | -25.1064
15.6941 | -34.8604 | -25.1141
15.7002 | -34.8609 | -25.1219
15.7063 | -34.8614 | -25.1299
15.7124 | -34.8619 | -25.1377
15.7185 | -34.8624 | -25.1454
15.7246 | -34.8629 | -25.1531
15.7307 | -34.8634 | -25.161
15.7368 | -34.8639 | -25.1691
15.7429 | -34.8644 | -25.1767
15.749 | -34.8649 | -25.1846
15.7551 | -34.8654 | -25.1923
15.7612 | -34.8659 | -25.1996
15.7673 | -34.8664 | -25.2072
15.7734 | -34.8668 | -25.2148
15.7795 | -34.8674 | -25.2229
15.7856 | -34.8679 | -25.2305
15.7917 | -34.8684 | -25.238
15.7978 | -34.8689 | -25.2457
15.8039 | -34.8694 | -25.2534
15.81 | -34.8699 | -25.2611
15.8161 | -34.8704 | -25.269
15.8222 | -34.8709 | -25.2768
15.8283 | -34.8714 | -25.2843
15.8345 | -34.8719 | -25.2921
15.8406 | -34.8724 | -25.2998
15.8467 | -34.873 | -25.3076
15.8528 | -34.8735 | -25.3153
15.8589 | -34.874 | -25.3229
15.865 | -34.8745 | -25.3305
15.8711 | -34.875 | -25.3385
15.8772 | -34.8755 | -25.3463
15.8833 | -34.876 | -25.3545
15.8894 | -34.8765 | -25.3622
15.8955 | -34.877 | -25.37
15.9016 | -34.8776 | -25.3776
15.9077 | -34.8781 | -25.3853
15.9138 | -34.8786 | -25.393
15.9199 | -34.8791 | -25.4009
15.926 | -34.8796 | -25.4089
15.9321 | -34.8801 | -25.4167
15.9382 | -34.8807 | -25.4243
15.9443 | -34.8812 | -25.4322
15.9504 | -34.8817 | -25.44
15.9565 | -34.8823 | -25.4477
15.9626 | -34.8828 | -25.4555
15.9687 | -34.8833 | -25.4629
15.9749 | -34.8838 | -25.4707
15.981 | -34.8844 | -25.4784
15.9871 | -34.8849 | -25.4861
15.9932 | -34.8854 | -25.494
15.9993 | -34.8859 | -25.5016
16.0054 | -34.8864 | -25.5089
16.0115 | -34.887 | -25.5165
16.0176 | -34.8875 | -25.5244
16.0237 | -34.888 | -25.5325
16.0298 | -34.8885 | -25.5396
16.0359 | -34.889 | -25.5472
16.042 | -34.8896 | -25.5548
16.0481 | -34.8901 | -25.5625
16.0542 | -34.8906 | -25.5705
16.0603 | -34.8912 | -25.578
16.0664 | -34.8916 | -25.5857
16.0725 | -34.8922 | -25.5931
16.0786 | -34.8928 | -25.601
16.0847 | -34.8933 | -25.6086
16.0908 | -34.8938 | -25.6163
16.0969 | -34.8943 | -25.6238
16.103 | -34.8948 | -25.6317
16.1091 | -34.8953 | -25.6393
16.1152 | -34.8958 | -25.6471
16.1214 | -34.8963 | -25.6549
16.1275 | -34.8969 | -25.6624
16.1336 | -34.8974 | -25.6702
16.1397 | -34.8979 | -25.6778
16.1458 | -34.8985 | -25.6854
16.1519 | -34.899 | -25.6933
16.158 | -34.8995 | -25.7011
16.1641 | -34.9 | -25.7089
16.1702 | -34.9005 | -25.7167
16.1763 | -34.901 | -25.7244
16.1824 | -34.9016 | -25.7317
16.1885 | -34.9021 | -25.7394
16.1946 | -34.9026 | -25.7471
16.2007 | -34.9031 | -25.7551
16.2068 | -34.9037 | -25.7627
16.2129 | -34.9041 | -25.7705
16.219 | -34.9046 | -25.7782
16.2251 | -34.9051 | -25.786
16.2312 | -34.9056 | -25.7939
16.2373 | -34.9061 | -25.8016
16.2434 | -34.9066 | -25.8092
16.2495 | -34.9072 | -25.8166
16.2556 | -34.9077 | -25.8244
16.2618 | -34.9082 | -25.8318
16.2679 | -34.9088 | -25.8395
16.274 | -34.9093 | -25.8473
16.2801 | -34.9098 | -25.8547
16.2862 | -34.9104 | -25.8622
16.2923 | -34.9109 | -25.87
16.2984 | -34.9114 | -25.8779
16.3045 | -34.9119 | -25.8854
16.3106 | -34.9125 | -25.8926
16.3167 | -34.913 | -25.9004
16.3228 | -34.9135 | -25.9082
16.3289 | -34.914 | -25.9157
16.335 | -34.9145 | -25.9235
16.3411 | -34.915 | -25.9309
16.3472 | -34.9155 | -25.9386
16.3533 | -34.9161 | -25.9463
16.3594 | -34.9166 | -25.9537
16.3655 | -34.9171 | -25.9615
16.3716 | -34.9176 | -25.969
16.3777 | -34.9181 | -25.9766
16.3838 | -34.9186 | -25.9842
16.3899 | -34.9191 | -25.9919
16.396 | -34.9196 | -25.9997
16.4021 | -34.9201 | -26.0073
16.4083 | -34.9207 | -26.0149
16.4144 | -34.9212 | -26.0223
16.4205 | -34.9217 | -26.0302
16.4266 | -34.9223 | -26.0376
16.4327 | -34.9228 | -26.045
16.4388 | -34.9233 | -26.0527
16.4449 | -34.9238 | -26.0603
16.451 | -34.9243 | -26.0682
16.4571 | -34.9248 | -26.0761
16.4632 | -34.9253 | -26.0834
16.4693 | -34.9258 | -26.0911
16.4754 | -34.9264 | -26.0989
16.4815 | -34.9269 | -26.1068
16.4876 | -34.9274 | -26.1143
16.4937 | -34.9279 | -26.1218
16.4998 | -34.9285 | -26.1294
16.5059 | -34.929 | -26.137
16.512 | -34.9295 | -26.145
16.5181 | -34.9301 | -26.1527
16.5242 | -34.9306 | -26.1604
16.5303 | -34.931 | -26.1678
16.5364 | -34.9315 | -26.1752
16.5425 | -34.932 | -26.1829
16.5487 | -34.9325 | -26.1905
16.5548 | -34.933 | -26.1981
16.5609 | -34.9336 | -26.2058
16.567 | -34.9341 | -26.2134
16.5731 | -34.9346 | -26.2212
16.5792 | -34.9351 | -26.2287
16.5853 | -34.9357 | -26.2362
16.5914 | -34.9362 | -26.244
16.5975 | -34.9367 | -26.2517
16.6036 | -34.9372 | -26.2592
16.6097 | -34.9377 | -26.2668
16.6158 | -34.9382 | -26.2746
16.6219 | -34.9388 | -26.2821
16.628 | -34.9392 | -26.2898
16.6341 | -34.9398 | -26.2971
16.6402 | -34.9403 | -26.3043
16.6463 | -34.9408 | -26.3119
16.6524 | -34.9413 | -26.3194
16.6585 | -34.9418 | -26.3272
16.6646 | -34.9424 | -26.3348
16.6707 | -34.9429 | -26.3422
16.6768 | -34.9434 | -26.3498
16.6829 | -34.9439 | -26.3574
16.689 | -34.9444 | -26.3651
16.6952 | -34.9449 | -26.3727
16.7013 | -34.9455 | -26.3802
16.7074 | -34.946 | -26.3877
16.7135 | -34.9465 | -26.3953
16.7196 | -34.947 | -26.4029
16.7257 | -34.9475 | -26.4103
16.7318 | -34.948 | -26.4179
16.7379 | -34.9485 | -26.4256
16.744 | -34.9491 | -26.4329
16.7501 | -34.9496 | -26.4405
16.7562 | -34.9501 | -26.448
16.7623 | -34.9506 | -26.4557
16.7684 | -34.9511 | -26.4634
16.7745 | -34.9517 | -26.4706
16.7806 | -34.9522 | -26.4782
16.7867 | -34.9527 | -26.4857
16.7928 | -34.9532 | -26.4934
16.7989 | -34.9536 | -26.5009
16.805 | -34.9541 | -26.5087
16.8111 | -34.9547 | -26.5163
16.8172 | -34.9552 | -26.5242
16.8233 | -34.9557 | -26.5319
16.8294 | -34.9563 | -26.5395
16.8356 | -34.9568 | -26.5473
16.8417 | -34.9573 | -26.5548
16.8478 | -34.9578 | -26.5623
16.8539 | -34.9583 | -26.5698
16.86 | -34.9588 | -26.5773
16.8661 | -34.9593 | -26.5848
16.8722 | -34.9599 | -26.5925
16.8783 | -34.9604 | -26.5997
16.8844 | -34.9609 | -26.6071
16.8905 | -34.9615 | -26.6148
16.8966 | -34.962 | -26.6223
16.9027 | -34.9625 | -26.6297
16.9088 | -34.9631 | -26.6375
16.9149 | -34.9636 | -26.6453
16.921 | -34.9641 | -26.6525
16.9271 | -34.9646 | -26.6605
16.9332 | -34.9652 | -26.6682
16.9393 | -34.9657 | -26.6759
16.9454 | -34.9661 | -26.6832
16.9515 | -34.9666 | -26.691
16.9576 | -34.9672 | -26.6986
16.9637 | -34.9677 | -26.7062
16.9698 | -34.9682 | -26.7135
16.9759 | -34.9688 | -26.7209
16.9821 | -34.9693 | -26.7285
16.9882 | -34.9698 | -26.7361
16.9943 | -34.9703 | -26.7438
17.0004 | -34.9708 | -26.7514
17.0065 | -34.9713 | -26.7589
17.0126 | -34.9719 | -26.7664
17.0187 | -34.9724 | -26.774
17.0248 | -34.9729 | -26.7813
17.0309 | -34.9734 | -26.7889
17.037 | -34.9739 | -26.7965
17.0431 | -34.9744 | -26.8037
17.0492 | -34.9749 | -26.8113
17.0553 | -34.9754 | -26.8189
17.0614 | -34.9759 | -26.8261
17.0675 | -34.9764 | -26.8339
17.0736 | -34.9769 | -26.8415
17.0797 | -34.9774 | -26.8487
17.0858 | -34.9779 | -26.8565
17.0919 | -34.9784 | -26.8642
17.098 | -34.9789 | -26.8716
17.1041 | -34.9794 | -26.8788
17.1102 | -34.9799 | -26.8867
17.1163 | -34.9805 | -26.8944
17.1225 | -34.981 | -26.9017
17.1286 | -34.9815 | -26.909
17.1347 | -34.982 | -26.9165
17.1408 | -34.9825 | -26.924
17.1469 | -34.983 | -26.9315
17.153 | -34.9836 | -26.939
17.1591 | -34.9841 | -26.9462
17.1652 | -34.9846 | -26.9535
17.1713 | -34.9851 | -26.961
17.1774 | -34.9856 | -26.9687
17.1835 | -34.9861 | -26.9764
17.1896 | -34.9866 | -26.9837
17.1957 | -34.9871 | -26.9913
17.2018 | -34.9876 | -26.9989
17.2079 | -34.9881 | -27.0062
17.214 | -34.9886 | -27.0137
17.2201 | -34.9891 | -27.0212
17.2262 | -34.9896 | -27.0289
17.2323 | -34.9901 | -27.0364
17.2384 | -34.9906 | -27.0441
17.2445 | -34.9911 | -27.0515
17.2506 | -34.9916 | -27.0591
17.2567 | -34.992 | -27.0665
17.2628 | -34.9926 | -27.0742
17.269 | -34.9931 | -27.0817
17.2751 | -34.9936 | -27.0892
17.2812 | -34.9941 | -27.0966
17.2873 | -34.9946 | -27.1041
17.2934 | -34.9951 | -27.1114
17.2995 | -34.9956 | -27.1188
17.3056 | -34.9961 | -27.126
17.3117 | -34.9965 | -27.1336
17.3178 | -34.9971 | -27.141
17.3239 | -34.9975 | -27.1484
17.33 | -34.998 | -27.1557
17.3361 | -34.9985 | -27.163
17.3422 | -34.999 | -27.1704
17.3483 | -34.9996 | -27.1781
17.3544 | -35.0001 | -27.1856
17.3605 | -35.0006 | -27.1931
17.3666 | -35.0011 | -27.2007
17.3727 | -35.0016 | -27.208
17.3788 | -35.0021 | -27.2152
17.3849 | -35.0026 | -27.2226
17.391 | -35.0031 | -27.2299
17.3971 | -35.0037 | -27.2374
17.4032 | -35.0041 | -27.2451
17.4094 | -35.0047 | -27.2525
17.4155 | -35.0051 | -27.2599
17.4216 | -35.0056 | -27.2673
17.4277 | -35.0061 | -27.275
17.4338 | -35.0067 | -27.2825
17.4399 | -35.0072 | -27.29
17.446 | -35.0077 | -27.2973
17.4521 | -35.0082 | -27.3045
17.4582 | -35.0087 | -27.3119
17.4643 | -35.0091 | -27.3194
17.4704 | -35.0096 | -27.3266
17.4765 | -35.0101 | -27.3341
17.4826 | -35.0106 | -27.3413
17.4887 | -35.0112 | -27.3491
17.4948 | -35.0117 | -27.3565
17.5009 | -35.0121 | -27.3637
17.507 | -35.0126 | -27.3713
17.5131 | -35.0131 | -27.3789
17.5192 | -35.0136 | -27.3863
17.5253 | -35.0141 | -27.3938
17.5314 | -35.0146 | -27.4012
17.5375 | -35.015 | -27.4087
17.5436 | -35.0155 | -27.4162
17.5497 | -35.016 | -27.4237
17.5559 | -35.0165 | -27.4309
17.562 | -35.017 | -27.4383
17.5681 | -35.0175 | -27.4458
17.5742 | -35.018 | -27.4535
17.5803 | -35.0185 | -27.4608
17.5864 | -35.019 | -27.468
17.5925 | -35.0195 | -27.4756
17.5986 | -35.02 | -27.4827
17.6047 | -35.0205 | -27.4905
17.6108 | -35.021 | -27.4977
17.6169 | -35.0215 | -27.5051
17.623 | -35.022 | -27.5122
17.6291 | -35.0225 | -27.5194
17.6352 | -35.023 | -27.5266
17.6413 | -35.0235 | -27.5342
17.6474 | -35.0239 | -27.5418
17.6535 | -35.0244 | -27.5492
17.6596 | -35.0249 | -27.5566
17.6657 | -35.0254 | -27.5639
17.6718 | -35.0259 | -27.5714
17.6779 | -35.0264 | -27.5788
17.684 | -35.0269 | -27.5864
17.6901 | -35.0274 | -27.5937
17.6963 | -35.0279 | -27.6009
17.7024 | -35.0284 | -27.6082
17.7085 | -35.0289 | -27.6154
17.7146 | -35.0294 | -27.6227
17.7207 | -35.0299 | -27.6299
17.7268 | -35.0304 | -27.6373
17.7329 | -35.0309 | -27.645
17.739 | -35.0313 | -27.6527
17.7451 | -35.0319 | -27.66
17.7512 | -35.0324 | -27.6675
17.7573 | -35.0328 | -27.675
17.7634 | -35.0333 | -27.6819
17.7695 | -35.0338 | -27.6893
17.7756 | -35.0343 | -27.6967
17.7817 | -35.0348 | -27.7038
17.7878 | -35.0353 | -27.7112
17.7939 | -35.0358 | -27.7184
17.8 | -35.0363 | -27.7257
17.8061 | -35.0368 | -27.7333
17.8122 | -35.0373 | -27.7406
17.8183 | -35.0378 | -27.7481
17.8244 | -35.0383 | -27.7554
17.8305 | -35.0388 | -27.7628
17.8366 | -35.0393 | -27.7698
17.8428 | -35.0398 | -27.7773
17.8489 | -35.0403 | -27.7846
17.855 | -35.0408 | -27.7918
17.8611 | -35.0413 | -27.7989
17.8672 | -35.0418 | -27.8061
17.8733 | -35.0423 | -27.8136
17.8794 | -35.0428 | -27.8211
17.8855 | -35.0433 | -27.8284
17.8916 | -35.0438 | -27.8356
17.8977 | -35.0443 | -27.8429
17.9038 | -35.0448 | -27.8501
17.9099 | -35.0453 | -27.8574
17.916 | -35.0457 | -27.8648
17.9221 | -35.0462 | -27.8721
17.9282 | -35.0467 | -27.8795
17.9343 | -35.0472 | -27.8869
17.9404 | -35.0477 | -27.894
17.9465 | -35.0482 | -27.9015
17.9526 | -35.0487 | -27.909
17.9587 | -35.0492 | -27.9163
17.9648 | -35.0496 | -27.9236
17.9709 | -35.0501 | -27.9308
17.977 | -35.0506 | -27.938
17.9832 | -35.0511 | -27.9455
17.9893 | -35.0516 | -27.953
17.9954 | -35.0521 | -27.9605
18.0015 | -35.0526 | -27.9679
18.0076 | -35.0531 | -27.9748
18.0137 | -35.0535 | -27.9824
18.0198 | -35.054 | -27.9896
18.0259 | -35.0545 | -27.997
18.032 | -35.0549 | -28.0041
18.0381 | -35.0554 | -28.0112
18.0442 | -35.0559 | -28.0187
18.0503 | -35.0564 | -28.0261
18.0564 | -35.0569 | -28.0335
18.0625 | -35.0573 | -28.0409
18.0686 | -35.0578 | -28.0482
18.0747 | -35.0583 | -28.0553
18.0808 | -35.0588 | -28.0623
18.0869 | -35.0593 | -28.0697
18.093 | -35.0598 | -28.0769
18.0991 | -35.0603 | -28.0844
18.1052 | -35.0608 | -28.092
18.1113 | -35.0612 | -28.099
18.1174 | -35.0617 | -28.1062
18.1236 | -35.0622 | -28.1135
18.1297 | -35.0627 | -28.1207
18.1358 | -35.0632 | -28.128
18.1419 | -35.0638 | -28.1353
18.148 | -35.0643 | -28.1426
18.1541 | -35.0648 | -28.1496
18.1602 | -35.0653 | -28.157
18.1663 | -35.0658 | -28.1641
18.1724 | -35.0663 | -28.1717
18.1785 | -35.0667 | -28.1791
18.1846 | -35.0672 | -28.1861
18.1907 | -35.0676 | -28.1937
18.1968 | -35.0681 | -28.2011
18.2029 | -35.0686 | -28.208
18.209 | -35.0691 | -28.2157
18.2151 | -35.0696 | -28.2232
18.2212 | -35.0701 | -28.2301
18.2273 | -35.0707 | -28.2373
18.2334 | -35.0711 | -28.2445
18.2395 | -35.0716 | -28.2516
18.2456 | -35.0721 | -28.2588
18.2517 | -35.0725 | -28.2661
18.2578 | -35.073 | -28.273
18.2639 | -35.0735 | -28.2806
18.2701 | -35.074 | -28.2877
18.2762 | -35.0745 | -28.2953
18.2823 | -35.075 | -28.3027
18.2884 | -35.0755 | -28.3101
18.2945 | -35.076 | -28.3171
18.3006 | -35.0765 | -28.3244
18.3067 | -35.0769 | -28.3315
18.3128 | -35.0774 | -28.3388
18.3189 | -35.0779 | -28.3464
18.325 | -35.0784 | -28.3534
18.3311 | -35.0789 | -28.3608
18.3372 | -35.0794 | -28.3678
18.3433 | -35.0799 | -28.375
18.3494 | -35.0803 | -28.3822
18.3555 | -35.0808 | -28.3891
18.3616 | -35.0813 | -28.3959
18.3677 | -35.0818 | -28.403
18.3738 | -35.0823 | -28.4102
18.3799 | -35.0828 | -28.4173
18.386 | -35.0832 | -28.4245
18.3921 | -35.0837 | -28.4315
18.3982 | -35.0842 | -28.4386
18.4043 | -35.0846 | -28.4458
18.4105 | -35.0851 | -28.4531
18.4166 | -35.0856 | -28.4603
18.4227 | -35.0861 | -28.4677
18.4288 | -35.0866 | -28.4751
18.4349 | -35.087 | -28.4822
18.441 | -35.0875 | -28.4894
18.4471 | -35.088 | -28.4966
18.4532 | -35.0885 | -28.5038
18.4593 | -35.089 | -28.5109
18.4654 | -35.0894 | -28.518
18.4715 | -35.0899 | -28.5253
18.4776 | -35.0904 | -28.5325
18.4837 | -35.0909 | -28.5398
18.4898 | -35.0914 | -28.547
18.4959 | -35.092 | -28.5544
18.502 | -35.0924 | -28.5616
18.5081 | -35.0929 | -28.569
18.5142 | -35.0934 | -28.5763
18.5203 | -35.0939 | -28.583
18.5264 | -35.0943 | -28.5901
18.5325 | -35.0949 | -28.5974
18.5386 | -35.0953 | -28.6049
18.5447 | -35.0958 | -28.612
18.5508 | -35.0963 | -28.6191
18.557 | -35.0968 | -28.6265
18.5631 | -35.0973 | -28.6337
18.5692 | -35.0977 | -28.6409
18.5753 | -35.0982 | -28.6478
18.5814 | -35.0987 | -28.6552
18.5875 | -35.0992 | -28.6627
18.5936 | -35.0996 | -28.6703
18.5997 | -35.1001 | -28.6774
18.6058 | -35.1006 | -28.6845
18.6119 | -35.1011 | -28.6919
18.618 | -35.1016 | -28.6991
18.6241 | -35.1021 | -28.7065
18.6302 | -35.1026 | -28.7139
18.6363 | -35.1031 | -28.7209
18.6424 | -35.1036 | -28.7281
18.6485 | -35.104 | -28.7354
18.6546 | -35.1045 | -28.7424
18.6607 | -35.105 | -28.7495
18.6668 | -35.1055 | -28.7567
18.6729 | -35.106 | -28.7639
18.679 | -35.1065 | -28.7709
18.6851 | -35.1069 | -28.778
18.6912 | -35.1074 | -28.7853
18.6974 | -35.1079 | -28.7925
18.7035 | -35.1084 | -28.7996
18.7096 | -35.1089 | -28.8068
18.7157 | -35.1094 | -28.8139
18.7218 | -35.1098 | -28.8211
18.7279 | -35.1104 | -28.8281
18.734 | -35.1109 | -28.8353
18.7401 | -35.1113 | -28.8422
18.7462 | -35.1118 | -28.8488
18.7523 | -35.1124 | -28.856
18.7584 | -35.1129 | -28.8633
18.7645 | -35.1134 | -28.8704
18.7706 | -35.1139 | -28.8776
18.7767 | -35.1143 | -28.8848
18.7828 | -35.1148 | -28.8915
18.7889 | -35.1153 | -28.8985
18.795 | -35.1157 | -28.9055
18.8011 | -35.1162 | -28.9127
18.8072 | -35.1167 | -28.9199
18.8133 | -35.1172 | -28.9271
18.8194 | -35.1176 | -28.9341
18.8255 | -35.1181 | -28.941
18.8316 | -35.1186 | -28.9482
18.8377 | -35.119 | -28.955
18.8439 | -35.1195 | -28.9618
18.85 | -35.12 | -28.969
18.8561 | -35.1205 | -28.976
18.8622 | -35.1209 | -28.9833
18.8683 | -35.1214 | -28.9903
18.8744 | -35.1219 | -28.9974
18.8805 | -35.1224 | -29.0047
18.8866 | -35.1229 | -29.0119
18.8927 | -35.1234 | -29.0187
18.8988 | -35.1239 | -29.0257
18.9049 | -35.1243 | -29.0327
18.911 | -35.1249 | -29.0396
18.9171 | -35.1254 | -29.0468
18.9232 | -35.1259 | -29.0538
18.9293 | -35.1264 | -29.0609
18.9354 | -35.1268 | -29.0679
18.9415 | -35.1273 | -29.0748
18.9476 | -35.1278 | -29.082
18.9537 | -35.1283 | -29.0894
18.9598 | -35.1288 | -29.0961
18.9659 | -35.1293 | -29.1031
18.972 | -35.1297 | -29.1101
18.9781 | -35.1302 | -29.1171
18.9843 | -35.1307 | -29.1241
18.9904 | -35.1311 | -29.1312
18.9965 | -35.1316 | -29.1382
19.0026 | -35.1321 | -29.1453
19.0087 | -35.1326 | -29.1525
19.0148 | -35.1331 | -29.1592
19.0209 | -35.1336 | -29.1667
19.027 | -35.1341 | -29.1739
19.0331 | -35.1345 | -29.181
19.0392 | -35.1351 | -29.1881
19.0453 | -35.1356 | -29.1952
19.0514 | -35.136 | -29.2023
19.0575 | -35.1365 | -29.2094
19.0636 | -35.137 | -29.2159
19.0697 | -35.1375 | -29.2232
19.0758 | -35.138 | -29.2301
19.0819 | -35.1385 | -29.2375
19.088 | -35.139 | -29.2443
19.0941 | -35.1395 | -29.251
19.1002 | -35.1399 | -29.2579
19.1063 | -35.1405 | -29.265
19.1124 | -35.141 | -29.2721
19.1185 | -35.1414 | -29.2795
19.1246 | -35.1419 | -29.2864
19.1308 | -35.1424 | -29.2932
19.1369 | -35.1429 | -29.3
19.143 | -35.1433 | -29.307
19.1491 | -35.1438 | -29.314
19.1552 | -35.1443 | -29.3213
19.1613 | -35.1448 | -29.3282
19.1674 | -35.1453 | -29.3353
19.1735 | -35.1458 | -29.3424
19.1796 | -35.1462 | -29.3492
19.1857 | -35.1467 | -29.3564
19.1918 | -35.1472 | -29.3636
19.1979 | -35.1477 | -29.3707
19.204 | -35.1482 | -29.3775
19.2101 | -35.1487 | -29.3845
19.2162 | -35.1491 | -29.3917
19.2223 | -35.1496 | -29.3982
19.2284 | -35.1501 | -29.4056
19.2345 | -35.1506 | -29.4125
19.2406 | -35.1511 | -29.4195
19.2467 | -35.1515 | -29.4265
19.2528 | -35.152 | -29.4336
19.2589 | -35.1524 | -29.4407
19.265 | -35.1529 | -29.4478
19.2712 | -35.1534 | -29.4549
19.2773 | -35.1539 | -29.462
19.2834 | -35.1544 | -29.4687
19.2895 | -35.1549 | -29.4759
19.2956 | -35.1554 | -29.4832
19.3017 | -35.1559 | -29.4899
19.3078 | -35.1564 | -29.497
19.3139 | -35.1569 | -29.504
19.32 | -35.1574 | -29.511
19.3261 | -35.1578 | -29.5179
19.3322 | -35.1583 | -29.5248
19.3383 | -35.1589 | -29.5315
19.3444 | -35.1594 | -29.5384
19.3505 | -35.1598 | -29.5454
19.3566 | -35.1602 | -29.5525
19.3627 | -35.1607 | -29.5597
19.3688 | -35.1612 | -29.567
19.3749 | -35.1617 | -29.5739
19.381 | -35.1622 | -29.5807
19.3871 | -35.1626 | -29.5878
19.3932 | -35.1631 | -29.5949
19.3993 | -35.1636 | -29.602
19.4054 | -35.1641 | -29.6092
19.4115 | -35.1646 | -29.6163
19.4177 | -35.165 | -29.6229
19.4238 | -35.1655 | -29.6297
19.4299 | -35.166 | -29.6363
19.436 | -35.1665 | -29.6438
19.4421 | -35.167 | -29.6503
19.4482 | -35.1675 | -29.6574
19.4543 | -35.168 | -29.664
19.4604 | -35.1685 | -29.6706
19.4665 | -35.169 | -29.6775
19.4726 | -35.1695 | -29.6842
19.4787 | -35.1699 | -29.6913
19.4848 | -35.1704 | -29.698
19.4909 | -35.1709 | -29.7049
19.497 | -35.1713 | -29.7117
19.5031 | -35.1718 | -29.7186
19.5092 | -35.1722 | -29.7254
19.5153 | -35.1727 | -29.7323
19.5214 | -35.1733 | -29.7392
19.5275 | -35.1737 | -29.7463
19.5336 | -35.1742 | -29.7529
19.5397 | -35.1746 | -29.7599
19.5458 | -35.1751 | -29.7668
19.5519 | -35.1756 | -29.7734
19.5581 | -35.1761 | -29.7798
19.5642 | -35.1766 | -29.7869
19.5703 | -35.177 | -29.7936
19.5764 | -35.1775 | -29.8004
19.5825 | -35.1779 | -29.807
19.5886 | -35.1783 | -29.8141
19.5947 | -35.1788 | -29.8208
19.6008 | -35.1792 | -29.8278
19.6069 | -35.1797 | -29.8347
19.613 | -35.1802 | -29.8418
19.6191 | -35.1807 | -29.8488
19.6252 | -35.1811 | -29.8558
19.6313 | -35.1816 | -29.8624
19.6374 | -35.182 | -29.8694
19.6435 | -35.1825 | -29.876
19.6496 | -35.183 | -29.8828
19.6557 | -35.1834 | -29.8895
19.6618 | -35.1839 | -29.8961
19.6679 | -35.1844 | -29.903
19.674 | -35.1849 | -29.9098
19.6801 | -35.1853 | -29.9165
19.6862 | -35.1858 | -29.9233
19.6923 | -35.1863 | -29.9301
19.6984 | -35.1868 | -29.9365
19.7046 | -35.1873 | -29.9434
19.7107 | -35.1877 | -29.9502
19.7168 | -35.1882 | -29.9574
19.7229 | -35.1886 | -29.9645
19.729 | -35.1891 | -29.9715
19.7351 | -35.1896 | -29.978
19.7412 | -35.1901 | -29.9848
19.7473 | -35.1905 | -29.9917
19.7534 | -35.191 | -29.9984
19.7595 | -35.1915 | -30.0055
19.7656 | -35.1919 | -30.0126
19.7717 | -35.1924 | -30.0191
19.7778 | -35.1928 | -30.0263
19.7839 | -35.1933 | -30.0328
19.79 | -35.1938 | -30.0397
19.7961 | -35.1943 | -30.0466
19.8022 | -35.1947 | -30.0533
19.8083 | -35.1952 | -30.0603
19.8144 | -35.1957 | -30.0671
19.8205 | -35.1962 | -30.0741
19.8266 | -35.1967 | -30.0812
19.8327 | -35.1972 | -30.0877
19.8388 | -35.1977 | -30.0946
19.845 | -35.1982 | -30.1013
19.8511 | -35.1986 | -30.1081
19.8572 | -35.1991 | -30.115
19.8633 | -35.1996 | -30.122
19.8694 | -35.2001 | -30.1291
19.8755 | -35.2006 | -30.1359
19.8816 | -35.2011 | -30.1425
19.8877 | -35.2016 | -30.1494
19.8938 | -35.202 | -30.1561
19.8999 | -35.2024 | -30.1628
19.906 | -35.2029 | -30.1699
19.9121 | -35.2034 | -30.1765
19.9182 | -35.2039 | -30.1831
19.9243 | -35.2044 | -30.1903
19.9304 | -35.2048 | -30.1977
19.9365 | -35.2053 | -30.2048
19.9426 | -35.2057 | -30.2115
19.9487 | -35.2062 | -30.2179
19.9548 | -35.2067 | -30.2249
19.9609 | -35.2073 | -30.2317
19.967 | -35.2078 | -30.2384
19.9731 | -35.2082 | -30.2454
19.9792 | -35.2087 | -30.2517
19.9853 | -35.2092 | -30.2588
19.9915 | -35.2097 | -30.2656
19.9976 | -35.2101 | -30.2725
20.0037 | -35.2106 | -30.2793
20.0098 | -35.211 | -30.2861
20.0159 | -35.2116 | -30.2926
20.022 | -35.2121 | -30.2998
20.0281 | -35.2127 | -30.3069
20.0342 | -35.2131 | -30.3134
20.0403 | -35.2136 | -30.3198
20.0464 | -35.2141 | -30.3272
20.0525 | -35.2145 | -30.3343
20.0586 | -35.215 | -30.341
20.0647 | -35.2155 | -30.3479
20.0708 | -35.216 | -30.355
20.0769 | -35.2164 | -30.3617
20.083 | -35.2168 | -30.3682
20.0891 | -35.2173 | -30.3752
20.0952 | -35.2178 | -30.3817
20.1013 | -35.2183 | -30.3883
20.1074 | -35.2187 | -30.3953
20.1135 | -35.2192 | -30.4023
20.1196 | -35.2196 | -30.409
20.1257 | -35.2201 | -30.4161
20.1319 | -35.2206 | -30.4229
20.138 | -35.221 | -30.4298
20.1441 | -35.2215 | -30.4366
20.1502 | -35.222 | -30.4436
20.1563 | -35.2225 | -30.4498
20.1624 | -35.223 | -30.4565
20.1685 | -35.2235 | -30.4632
20.1746 | -35.2239 | -30.47
20.1807 | -35.2245 | -30.4769
20.1868 | -35.225 | -30.484
20.1929 | -35.2254 | -30.4904
20.199 | -35.2259 | -30.4972
20.2051 | -35.2263 | -30.5038
20.2112 | -35.2268 | -30.5106
20.2173 | -35.2272 | -30.5177
20.2234 | -35.2277 | -30.5246
20.2295 | -35.2282 | -30.5311
20.2356 | -35.2287 | -30.5379
20.2417 | -35.2292 | -30.5447
20.2478 | -35.2296 | -30.5517
20.2539 | -35.2301 | -30.559
20.26 | -35.2305 | -30.5658
20.2661 | -35.231 | -30.5727
20.2723 | -35.2315 | -30.5794
20.2784 | -35.2319 | -30.5861
20.2845 | -35.2324 | -30.5928
20.2906 | -35.2329 | -30.5995
20.2967 | -35.2334 | -30.6062
20.3028 | -35.2339 | -30.6128
20.3089 | -35.2343 | -30.6196
20.315 | -35.2348 | -30.6262
20.3211 | -35.2352 | -30.6328
20.3272 | -35.2357 | -30.6394
20.3333 | -35.2362 | -30.6461
20.3394 | -35.2367 | -30.6527
20.3455 | -35.2372 | -30.6595
20.3516 | -35.2376 | -30.6663
20.3577 | -35.2381 | -30.6729
20.3638 | -35.2386 | -30.6798
20.3699 | -35.2391 | -30.6862
20.376 | -35.2396 | -30.693
20.3821 | -35.24 | -30.6997
20.3882 | -35.2405 | -30.7064
20.3943 | -35.241 | -30.7133
20.4004 | -35.2415 | -30.7198
20.4065 | -35.2419 | -30.7264
20.4126 | -35.2424 | -30.7331
20.4188 | -35.2429 | -30.74
20.4249 | -35.2434 | -30.7469
20.431 | -35.2438 | -30.7536
20.4371 | -35.2443 | -30.7597
20.4432 | -35.2447 | -30.7666
20.4493 | -35.2452 | -30.7732
20.4554 | -35.2457 | -30.7797
20.4615 | -35.2462 | -30.7866
20.4676 | -35.2466 | -30.7934
20.4737 | -35.2471 | -30.7997
20.4798 | -35.2476 | -30.8065
20.4859 | -35.2481 | -30.8131
20.492 | -35.2486 | -30.8199
20.4981 | -35.249 | -30.8264
20.5042 | -35.2495 | -30.8331
20.5103 | -35.2499 | -30.8395
20.5164 | -35.2504 | -30.8462
20.5225 | -35.2508 | -30.8529
20.5286 | -35.2513 | -30.86
20.5347 | -35.2517 | -30.8666
20.5408 | -35.2522 | -30.8734
20.5469 | -35.2527 | -30.8799
20.553 | -35.2532 | -30.8867
20.5592 | -35.2536 | -30.8932
20.5653 | -35.2541 | -30.9
20.5714 | -35.2546 | -30.9065
20.5775 | -35.2551 | -30.9131
20.5836 | -35.2555 | -30.9199
20.5897 | -35.256 | -30.9266
20.5958 | -35.2565 | -30.9333
20.6019 | -35.257 | -30.9399
20.608 | -35.2574 | -30.9465
20.6141 | -35.2579 | -30.9528
20.6202 | -35.2583 | -30.9595
20.6263 | -35.2588 | -30.9659
20.6324 | -35.2592 | -30.9725
20.6385 | -35.2597 | -30.979
20.6446 | -35.2601 | -30.9859
20.6507 | -35.2606 | -30.9921
20.6568 | -35.261 | -30.9987
20.6629 | -35.2615 | -31.0053
20.669 | -35.262 | -31.012
20.6751 | -35.2624 | -31.0186
20.6812 | -35.2629 | -31.0248
20.6873 | -35.2634 | -31.0316
20.6934 | -35.2639 | -31.0383
20.6995 | -35.2643 | -31.0447
20.7057 | -35.2648 | -31.0512
20.7118 | -35.2653 | -31.0576
20.7179 | -35.2658 | -31.0638
20.724 | -35.2662 | -31.0701
20.7301 | -35.2666 | -31.0769
20.7362 | -35.2671 | -31.0837
20.7423 | -35.2676 | -31.0902
20.7484 | -35.268 | -31.0968
20.7545 | -35.2685 | -31.1031
20.7606 | -35.269 | -31.1099
20.7667 | -35.2695 | -31.1166
20.7728 | -35.27 | -31.1231
20.7789 | -35.2705 | -31.1297
20.785 | -35.271 | -31.1364
20.7911 | -35.2714 | -31.1429
20.7972 | -35.2719 | -31.1495
20.8033 | -35.2723 | -31.1558
20.8094 | -35.2728 | -31.1622
20.8155 | -35.2732 | -31.1689
20.8216 | -35.2737 | -31.1755
20.8277 | -35.2741 | -31.1817
20.8338 | -35.2745 | -31.1883
20.8399 | -35.275 | -31.1947
20.8461 | -35.2755 | -31.2012
20.8522 | -35.276 | -31.208
20.8583 | -35.2765 | -31.2146
20.8644 | -35.277 | -31.2211
20.8705 | -35.2775 | -31.2278
20.8766 | -35.2779 | -31.2346
20.8827 | -35.2784 | -31.2411
20.8888 | -35.2789 | -31.2475
20.8949 | -35.2793 | -31.254
20.901 | -35.2798 | -31.2606
20.9071 | -35.2803 | -31.2673
20.9132 | -35.2807 | -31.2739
20.9193 | -35.2812 | -31.2806
20.9254 | -35.2817 | -31.287
20.9315 | -35.2822 | -31.2931
20.9376 | -35.2826 | -31.2995
20.9437 | -35.2831 | -31.3064
20.9498 | -35.2836 | -31.313
20.9559 | -35.2841 | -31.3193
20.962 | -35.2846 | -31.3261
20.9681 | -35.285 | -31.3326
20.9742 | -35.2855 | -31.339
20.9803 | -35.286 | -31.3454
20.9864 | -35.2864 | -31.352
20.9926 | -35.2869 | -31.3587
20.9987 | -35.2874 | -31.365
21.0048 | -35.2878 | -31.3715
21.0109 | -35.2883 | -31.3781
21.017 | -35.2887 | -31.3845
21.0231 | -35.2892 | -31.3909
21.0292 | -35.2897 | -31.3973
21.0353 | -35.2901 | -31.4034
21.0414 | -35.2906 | -31.41
21.0475 | -35.2911 | -31.4166
21.0536 | -35.2916 | -31.4231
21.0597 | -35.292 | -31.4294
21.0658 | -35.2925 | -31.4359
21.0719 | -35.293 | -31.4425
21.078 | -35.2935 | -31.4493
21.0841 | -35.294 | -31.4555
21.0902 | -35.2944 | -31.4623
21.0963 | -35.2948 | -31.4685
21.1024 | -35.2953 | -31.4748
21.1085 | -35.2957 | -31.4815
21.1146 | -35.2962 | -31.4885
21.1207 | -35.2967 | -31.495
21.1268 | -35.2971 | -31.5016
21.133 | -35.2976 | -31.5079
21.1391 | -35.298 | -31.5145
21.1452 | -35.2985 | -31.5211
21.1513 | -35.2989 | -31.5276
21.1574 | -35.2994 | -31.5344
21.1635 | -35.2998 | -31.5409
21.1696 | -35.3003 | -31.5473
21.1757 | -35.3008 | -31.5533
21.1818 | -35.3012 | -31.5601
21.1879 | -35.3017 | -31.5664
21.194 | -35.3022 | -31.5729
21.2001 | -35.3026 | -31.5792
21.2062 | -35.3031 | -31.5856
21.2123 | -35.3036 | -31.5921
21.2184 | -35.304 | -31.5988
21.2245 | -35.3045 | -31.6053
21.2306 | -35.3049 | -31.6114
21.2367 | -35.3054 | -31.6178
21.2428 | -35.3059 | -31.6241
21.2489 | -35.3063 | -31.6305
21.255 | -35.3068 | -31.6371
21.2611 | -35.3073 | -31.6436
21.2672 | -35.3077 | -31.6502
21.2733 | -35.3082 | -31.6565
21.2795 | -35.3087 | -31.6628
21.2856 | -35.3091 | -31.6692
21.2917 | -35.3096 | -31.6757
21.2978 | -35.31 | -31.6825
21.3039 | -35.3104 | -31.6888
21.31 | -35.3109 | -31.695
21.3161 | -35.3114 | -31.7014
21.3222 | -35.3119 | -31.7079
21.3283 | -35.3124 | -31.7145
21.3344 | -35.3128 | -31.721
21.3405 | -35.3132 | -31.7276
21.3466 | -35.3137 | -31.7341
21.3527 | -35.3142 | -31.7403
21.3588 | -35.3146 | -31.7469
21.3649 | -35.3151 | -31.7533
21.371 | -35.3156 | -31.7594
21.3771 | -35.316 | -31.7656
21.3832 | -35.3165 | -31.7719
21.3893 | -35.3169 | -31.7779
21.3954 | -35.3173 | -31.7842
21.4015 | -35.3178 | -31.7904
21.4076 | -35.3183 | -31.797
21.4137 | -35.3187 | -31.8036
21.4199 | -35.3192 | -31.8102
21.426 | -35.3197 | -31.8167
21.4321 | -35.3202 | -31.8231
21.4382 | -35.3206 | -31.8294
21.4443 | -35.3211 | -31.8356
21.4504 | -35.3215 | -31.8421
21.4565 | -35.322 | -31.8485
21.4626 | -35.3225 | -31.855
21.4687 | -35.323 | -31.8615
21.4748 | -35.3234 | -31.8679
21.4809 | -35.3238 | -31.8743
21.487 | -35.3243 | -31.8807
21.4931 | -35.3247 | -31.8869
21.4992 | -35.3252 | -31.8929
21.5053 | -35.3256 | -31.8995
21.5114 | -35.3261 | -31.9057
21.5175 | -35.3265 | -31.9119
21.5236 | -35.3269 | -31.9183
21.5297 | -35.3274 | -31.9245
21.5358 | -35.3278 | -31.9309
21.5419 | -35.3283 | -31.9376
21.548 | -35.3287 | -31.9441
21.5541 | -35.3291 | -31.9506
21.5602 | -35.3296 | -31.957
21.5664 | -35.3301 | -31.9633
21.5725 | -35.3305 | -31.9698
21.5786 | -35.3309 | -31.9758
21.5847 | -35.3313 | -31.9823
21.5908 | -35.3318 | -31.9886
21.5969 | -35.3323 | -31.9948
21.603 | -35.3328 | -32.001
21.6091 | -35.3333 | -32.0076
21.6152 | -35.3338 | -32.0141
21.6213 | -35.3342 | -32.0206
21.6274 | -35.3348 | -32.027
21.6335 | -35.3352 | -32.0332
21.6396 | -35.3356 | -32.0397
21.6457 | -35.336 | -32.0462
21.6518 | -35.3364 | -32.0523
21.6579 | -35.3369 | -32.0583
21.664 | -35.3374 | -32.0651
21.6701 | -35.3379 | -32.0713
21.6762 | -35.3384 | -32.0778
21.6823 | -35.3388 | -32.084
21.6884 | -35.3392 | -32.0905
21.6945 | -35.3397 | -32.0969
21.7006 | -35.3401 | -32.1035
21.7068 | -35.3406 | -32.1095
21.7129 | -35.3411 | -32.1157
21.719 | -35.3416 | -32.1222
21.7251 | -35.342 | -32.1284
21.7312 | -35.3424 | -32.1349
21.7373 | -35.3429 | -32.141
21.7434 | -35.3433 | -32.1476
21.7495 | -35.3438 | -32.1539
21.7556 | -35.3442 | -32.1606
21.7617 | -35.3447 | -32.1671
21.7678 | -35.3452 | -32.1736
21.7739 | -35.3457 | -32.1797
21.78 | -35.3461 | -32.1861
21.7861 | -35.3466 | -32.1924
21.7922 | -35.347 | -32.1989
21.7983 | -35.3475 | -32.2054
21.8044 | -35.3479 | -32.2119
21.8105 | -35.3484 | -32.2182
21.8166 | -35.3489 | -32.2246
21.8227 | -35.3494 | -32.2309
21.8288 | -35.3499 | -32.237
21.8349 | -35.3503 | -32.2428
21.841 | -35.3508 | -32.2492
21.8471 | -35.3512 | -32.2556
21.8533 | -35.3516 | -32.2617
21.8594 | -35.3521 | -32.2685
21.8655 | -35.3525 | -32.275
21.8716 | -35.3531 | -32.2814
21.8777 | -35.3535 | -32.2877
21.8838 | -35.354 | -32.2941
21.8899 | -35.3545 | -32.3004
21.896 | -35.3549 | -32.3068
21.9021 | -35.3554 | -32.3133
21.9082 | -35.3558 | -32.3196
21.9143 | -35.3563 | -32.3254
21.9204 | -35.3567 | -32.3317
21.9265 | -35.3572 | -32.338
21.9326 | -35.3576 | -32.3442
21.9387 | -35.3581 | -32.3506
21.9448 | -35.3585 | -32.3568
21.9509 | -35.359 | -32.3627
21.957 | -35.3595 | -32.3687
21.9631 | -35.3599 | -32.3752
21.9692 | -35.3604 | -32.3814
21.9753 | -35.3608 | -32.3878
21.9814 | -35.3613 | -32.3943
21.9875 | -35.3617 | -32.4001
21.9937 | -35.3622 | -32.4065
21.9998 | -35.3627 | -32.4132
22.0059 | -35.3631 | -32.4194
22.012 | -35.3636 | -32.4257
22.0181 | -35.3641 | -32.4321
22.0242 | -35.3645 | -32.4385
22.0303 | -35.365 | -32.445
22.0364 | -35.3654 | -32.4515
22.0425 | -35.3659 | -32.4579
22.0486 | -35.3663 | -32.4643
22.0547 | -35.3667 | -32.4702
22.0608 | -35.3672 | -32.4763
22.0669 | -35.3676 | -32.4824
22.073 | -35.3681 | -32.4888
22.0791 | -35.3686 | -32.4951
22.0852 | -35.369 | -32.5013
22.0913 | -35.3695 | -32.5078
22.0974 | -35.3699 | -32.514
22.1035 | -35.3703 | -32.5203
22.1096 | -35.3708 | -32.5268
22.1157 | -35.3712 | -32.5329
22.1218 | -35.3717 | -32.5394
22.1279 | -35.3722 | -32.5452
22.134 | -35.3726 | -32.5514
22.1402 | -35.373 | -32.5575
22.1463 | -35.3735 | -32.5635
22.1524 | -35.374 | -32.57
22.1585 | -35.3743 | -32.5762
22.1646 | -35.3748 | -32.582
22.1707 | -35.3752 | -32.5886
22.1768 | -35.3756 | -32.5948
22.1829 | -35.3761 | -32.601
22.189 | -35.3765 | -32.6072
22.1951 | -35.377 | -32.613
22.2012 | -35.3774 | -32.6197
22.2073 | -35.3779 | -32.6259
22.2134 | -35.3784 | -32.6321
22.2195 | -35.3788 | -32.6382
22.2256 | -35.3793 | -32.6444
22.2317 | -35.3797 | -32.6507
22.2378 | -35.3802 | -32.6569
22.2439 | -35.3807 | -32.6635
22.25 | -35.3811 | -32.6698
22.2561 | -35.3816 | -32.6761
22.2622 | -35.3821 | -32.6825
22.2683 | -35.3825 | -32.6883
22.2744 | -35.3829 | -32.6948
22.2806 | -35.3834 | -32.7009
22.2867 | -35.3839 | -32.7069
22.2928 | -35.3843 | -32.7132
22.2989 | -35.3847 | -32.7195
22.305 | -35.3852 | -32.7256
22.3111 | -35.3856 | -32.7318
22.3172 | -35.3861 | -32.7383
22.3233 | -35.3866 | -32.7446
22.3294 | -35.3871 | -32.7509
22.3355 | -35.3876 | -32.7571
22.3416 | -35.3881 | -32.7634
22.3477 | -35.3885 | -32.7698
22.3538 | -35.3889 | -32.7757
22.3599 | -35.3893 | -32.7818
22.366 | -35.3897 | -32.7882
22.3721 | -35.3902 | -32.794
22.3782 | -35.3907 | -32.8002
22.3843 | -35.3911 | -32.8067
22.3904 | -35.3916 | -32.8131
22.3965 | -35.3921 | -32.8194
22.4026 | -35.3925 | -32.8256
22.4087 | -35.3929 | -32.8316
22.4148 | -35.3933 | -32.8378
22.4209 | -35.3938 | -32.8442
22.4271 | -35.3942 | -32.8503
22.4332 | -35.3947 | -32.8561
22.4393 | -35.3951 | -32.8621
22.4454 | -35.3955 | -32.8681
22.4515 | -35.3959 | -32.8744
22.4576 | -35.3964 | -32.8806
22.4637 | -35.3968 | -32.887
22.4698 | -35.3973 | -32.8932
22.4759 | -35.3977 | -32.8993
22.482 | -35.3981 | -32.9058
22.4881 | -35.3985 | -32.9121
22.4942 | -35.399 | -32.9183
22.5003 | -35.3994 | -32.9245
22.5064 | -35.3998 | -32.9308
22.5125 | -35.4002 | -32.9367
22.5186 | -35.4007 | -32.9428
22.5247 | -35.4011 | -32.949
22.5308 | -35.4016 | -32.9552
22.5369 | -35.402 | -32.9617
22.543 | -35.4025 | -32.968
22.5491 | -35.4028 | -32.9739
22.5552 | -35.4033 | -32.9802
22.5613 | -35.4037 | -32.9868
22.5675 | -35.4042 | -32.9933
22.5736 | -35.4046 | -32.9995
22.5797 | -35.405 | -33.0055
22.5858 | -35.4054 | -33.0117
22.5919 | -35.4059 | -33.0179
22.598 | -35.4063 | -33.0239
22.6041 | -35.4067 | -33.0301
22.6102 | -35.4072 | -33.0362
22.6163 | -35.4076 | -33.042
22.6224 | -35.4081 | -33.0481
22.6285 | -35.4086 | -33.0541
22.6346 | -35.409 | -33.0601
22.6407 | -35.4094 | -33.0666
22.6468 | -35.4099 | -33.0729
22.6529 | -35.4103 | -33.0791
22.659 | -35.4107 | -33.0853
22.6651 | -35.4112 | -33.0916
22.6712 | -35.4116 | -33.0978
22.6773 | -35.412 | -33.1041
22.6834 | -35.4124 | -33.1105
22.6895 | -35.4129 | -33.116
22.6956 | -35.4133 | -33.1221
22.7017 | -35.4138 | -33.1284
22.7079 | -35.4143 | -33.1349
22.714 | -35.4147 | -33.1411
22.7201 | -35.4151 | -33.1474
22.7262 | -35.4156 | -33.1533
22.7323 | -35.416 | -33.1594
22.7384 | -35.4164 | -33.1658
22.7445 | -35.4168 | -33.1719
22.7506 | -35.4172 | -33.178
22.7567 | -35.4177 | -33.1843
22.7628 | -35.4181 | -33.1903
22.7689 | -35.4185 | -33.1967
22.775 | -35.4189 | -33.203
22.7811 | -35.4194 | -33.2092
22.7872 | -35.4198 | -33.2154
22.7933 | -35.4203 | -33.2215
22.7994 | -35.4207 | -33.2277
22.8055 | -35.4212 | -33.2343
22.8116 | -35.4216 | -33.2406
22.8177 | -35.422 | -33.2469
22.8238 | -35.4224 | -33.2532
22.8299 | -35.4228 | -33.2589
22.836 | -35.4233 | -33.2653
22.8421 | -35.4237 | -33.2715
22.8482 | -35.4242 | -33.2776
22.8544 | -35.4246 | -33.2838
22.8605 | -35.4251 | -33.2898
22.8666 | -35.4255 | -33.296
22.8727 | -35.426 | -33.3023
22.8788 | -35.4264 | -33.3083
22.8849 | -35.4269 | -33.3141
22.891 | -35.4273 | -33.3203
22.8971 | -35.4277 | -33.3265
22.9032 | -35.4282 | -33.3326
22.9093 | -35.4286 | -33.3386
22.9154 | -35.4291 | -33.3448
22.9215 | -35.4295 | -33.3509
22.9276 | -35.4299 | -33.3571
22.9337 | -35.4303 | -33.3629
22.9398 | -35.4308 | -33.3693
22.9459 | -35.4312 | -33.3754
22.952 | -35.4317 | -33.3816
22.9581 | -35.432 | -33.3879
22.9642 | -35.4325 | -33.3941
22.9703 | -35.4329 | -33.4001
22.9764 | -35.4334 | -33.4061
22.9825 | -35.4338 | -33.4121
22.9886 | -35.4342 | -33.4182
22.9948 | -35.4347 | -33.4246
23.0009 | -35.4351 | -33.4308
23.007 | -35.4355 | -33.4369
23.0131 | -35.436 | -33.4429
23.0192 | -35.4364 | -33.4488
23.0253 | -35.4369 | -33.4549
23.0314 | -35.4373 | -33.4609
23.0375 | -35.4378 | -33.4672
23.0436 | -35.4382 | -33.4731
23.0497 | -35.4387 | -33.4792
23.0558 | -35.4391 | -33.4853
23.0619 | -35.4396 | -33.4913
23.068 | -35.4401 | -33.4974
23.0741 | -35.4405 | -33.5032
23.0802 | -35.4409 | -33.5094
23.0863 | -35.4413 | -33.5156
23.0924 | -35.4417 | -33.5217
23.0985 | -35.4421 | -33.5277
23.1046 | -35.4426 | -33.5338
23.1107 | -35.4429 | -33.5402
23.1168 | -35.4434 | -33.5465
23.1229 | -35.4438 | -33.5525
23.129 | -35.4442 | -33.5588
23.1351 | -35.4447 | -33.565
23.1413 | -35.4451 | -33.5709
23.1474 | -35.4456 | -33.5767
23.1535 | -35.446 | -33.5828
23.1596 | -35.4464 | -33.5886
23.1657 | -35.4468 | -33.5946
23.1718 | -35.4472 | -33.6006
23.1779 | -35.4476 | -33.6069
23.184 | -35.4481 | -33.6125
23.1901 | -35.4485 | -33.6187
23.1962 | -35.449 | -33.6248
23.2023 | -35.4494 | -33.6309
23.2084 | -35.4499 | -33.6371
23.2145 | -35.4503 | -33.6433
23.2206 | -35.4507 | -33.6495
23.2267 | -35.4511 | -33.6556
23.2328 | -35.4515 | -33.6616
23.2389 | -35.452 | -33.6679
23.245 | -35.4524 | -33.674
23.2511 | -35.4528 | -33.6799
23.2572 | -35.4533 | -33.6857
23.2633 | -35.4537 | -33.6918
23.2694 | -35.4542 | -33.6979
23.2755 | -35.4546 | -33.7041
23.2817 | -35.455 | -33.7101
23.2878 | -35.4554 | -33.716
23.2939 | -35.4559 | -33.7221
23.3 | -35.4563 | -33.7283
23.3061 | -35.4568 | -33.7343
23.3122 | -35.4572 | -33.7402
23.3183 | -35.4576 | -33.7464
23.3244 | -35.458 | -33.7525
23.3305 | -35.4585 | -33.7587
23.3366 | -35.4588 | -33.7649
23.3427 | -35.4593 | -33.7707
23.3488 | -35.4597 | -33.7769
23.3549 | -35.4601 | -33.7832
23.361 | -35.4606 | -33.7891
23.3671 | -35.4611 | -33.7952
23.3732 | -35.4615 | -33.8011
23.3793 | -35.4619 | -33.8071
23.3854 | -35.4624 | -33.8134
23.3915 | -35.4628 | -33.8193
23.3976 | -35.4632 | -33.8255
23.4037 | -35.4636 | -33.8313
23.4098 | -35.4641 | -33.8374
23.4159 | -35.4645 | -33.8433
23.422 | -35.4649 | -33.8493
23.4282 | -35.4653 | -33.8551
23.4343 | -35.4658 | -33.8611
23.4404 | -35.4662 | -33.8673
23.4465 | -35.4665 | -33.8733
23.4526 | -35.467 | -33.8795
23.4587 | -35.4674 | -33.8854
23.4648 | -35.4677 | -33.8913
23.4709 | -35.4681 | -33.8973
23.477 | -35.4685 | -33.9036
23.4831 | -35.4689 | -33.9096
23.4892 | -35.4694 | -33.9158
23.4953 | -35.4698 | -33.9219
23.5014 | -35.4702 | -33.9277
23.5075 | -35.4707 | -33.9338
23.5136 | -35.4711 | -33.9397
23.5197 | -35.4715 | -33.9459
23.5258 | -35.4719 | -33.9519
23.5319 | -35.4724 | -33.9576
23.538 | -35.4728 | -33.9635
23.5441 | -35.4732 | -33.9695
23.5502 | -35.4736 | -33.9756
23.5563 | -35.474 | -33.9817
23.5624 | -35.4744 | -33.9879
23.5686 | -35.4748 | -33.9938
23.5747 | -35.4752 | -33.9996
23.5808 | -35.4757 | -34.0056
23.5869 | -35.4761 | -34.0115
23.593 | -35.4765 | -34.0176
23.5991 | -35.4769 | -34.0236
23.6052 | -35.4773 | -34.0292
23.6113 | -35.4778 | -34.0352
23.6174 | -35.4782 | -34.0413
23.6235 | -35.4786 | -34.0473
23.6296 | -35.479 | -34.0535
23.6357 | -35.4794 | -34.0598
23.6418 | -35.4798 | -34.0654
23.6479 | -35.4803 | -34.0717
23.654 | -35.4807 | -34.0776
23.6601 | -35.4811 | -34.0836
23.6662 | -35.4815 | -34.0897
23.6723 | -35.4819 | -34.0956
23.6784 | -35.4823 | -34.1015
23.6845 | -35.4827 | -34.1075
23.6906 | -35.4831 | -34.1133
23.6967 | -35.4835 | -34.1192
23.7028 | -35.4839 | -34.1253
23.7089 | -35.4843 | -34.1312
23.7151 | -35.4847 | -34.1371
23.7212 | -35.4851 | -34.1432
23.7273 | -35.4855 | -34.1492
23.7334 | -35.486 | -34.1552
23.7395 | -35.4864 | -34.161
23.7456 | -35.4868 | -34.1669
23.7517 | -35.4872 | -34.1727
23.7578 | -35.4876 | -34.1788
23.7639 | -35.488 | -34.1848
23.77 | -35.4884 | -34.1907
23.7761 | -35.4888 | -34.1965
23.7822 | -35.4892 | -34.2025
23.7883 | -35.4896 | -34.2085
23.7944 | -35.49 | -34.2148
23.8005 | -35.4904 | -34.221
23.8066 | -35.4909 | -34.2271
23.8127 | -35.4913 | -34.2329
23.8188 | -35.4917 | -34.2389
23.8249 | -35.4921 | -34.2451
23.831 | -35.4925 | -34.2514
23.8371 | -35.4929 | -34.2572
23.8432 | -35.4934 | -34.263
23.8493 | -35.4938 | -34.2687
23.8555 | -35.4943 | -34.2747
23.8616 | -35.4947 | -34.2807
23.8677 | -35.4951 | -34.2865
23.8738 | -35.4955 | -34.2923
23.8799 | -35.496 | -34.2982
23.886 | -35.4963 | -34.3041
23.8921 | -35.4967 | -34.3102
23.8982 | -35.4972 | -34.3161
23.9043 | -35.4976 | -34.3223
23.9104 | -35.4979 | -34.3282
23.9165 | -35.4983 | -34.334
23.9226 | -35.4987 | -34.3397
23.9287 | -35.4991 | -34.3457
23.9348 | -35.4995 | -34.3515
23.9409 | -35.4999 | -34.3575
23.947 | -35.5004 | -34.3634
23.9531 | -35.5008 | -34.369
23.9592 | -35.5012 | -34.3751
23.9653 | -35.5016 | -34.381
23.9714 | -35.502 | -34.387
23.9775 | -35.5024 | -34.393
23.9836 | -35.5028 | -34.3991
23.9897 | -35.5032 | -34.405
23.9958 | -35.5036 | -34.4111
24.002 | -35.504 | -34.4169
24.0081 | -35.5044 | -34.4228
24.0142 | -35.5048 | -34.4288
24.0203 | -35.5053 | -34.4347
24.0264 | -35.5056 | -34.4405
24.0325 | -35.5061 | -34.4464
24.0386 | -35.5064 | -34.4525
24.0447 | -35.5068 | -34.4584
24.0508 | -35.5073 | -34.4643
24.0569 | -35.5077 | -34.4703
24.063 | -35.5081 | -34.476
24.0691 | -35.5085 | -34.4819
24.0752 | -35.5089 | -34.4881
24.0813 | -35.5093 | -34.4937
24.0874 | -35.5097 | -34.4996
24.0935 | -35.5102 | -34.5056
24.0996 | -35.5106 | -34.5111
24.1057 | -35.511 | -34.5169
24.1118 | -35.5114 | -34.5228
24.1179 | -35.5117 | -34.5289
24.124 | -35.5122 | -34.5348
24.1301 | -35.5126 | -34.5405
24.1362 | -35.513 | -34.5465
24.1424 | -35.5134 | -34.5523
24.1485 | -35.5138 | -34.5582
24.1546 | -35.5142 | -34.5642
24.1607 | -35.5146 | -34.5701
24.1668 | -35.515 | -34.5756
24.1729 | -35.5153 | -34.5814
24.179 | -35.5157 | -34.5874
24.1851 | -35.5161 | -34.5934
24.1912 | -35.5166 | -34.5996
24.1973 | -35.517 | -34.6056
24.2034 | -35.5174 | -34.611
24.2095 | -35.5178 | -34.617
24.2156 | -35.5182 | -34.6231
24.2217 | -35.5187 | -34.629
24.2278 | -35.5191 | -34.6349
24.2339 | -35.5195 | -34.6408
24.24 | -35.5199 | -34.6465
24.2461 | -35.5203 | -34.6525
24.2522 | -35.5207 | -34.6583
24.2583 | -35.5212 | -34.6644
24.2644 | -35.5216 | -34.6703
24.2705 | -35.522 | -34.6759
24.2766 | -35.5224 | -34.6817
24.2827 | -35.5228 | -34.6874
24.2889 | -35.5232 | -34.6933
24.295 | -35.5237 | -34.6992
24.3011 | -35.524 | -34.7052
24.3072 | -35.5245 | -34.7108
24.3133 | -35.5249 | -34.7164
24.3194 | -35.5252 | -34.7222
24.3255 | -35.5256 | -34.7282
24.3316 | -35.5261 | -34.734
24.3377 | -35.5265 | -34.7398
24.3438 | -35.5268 | -34.7456
24.3499 | -35.5272 | -34.7514
24.356 | -35.5276 | -34.7571
24.3621 | -35.528 | -34.7631
24.3682 | -35.5283 | -34.7689
24.3743 | -35.5287 | -34.775
24.3804 | -35.5291 | -34.7803
24.3865 | -35.5295 | -34.7863
24.3926 | -35.5299 | -34.7921
24.3987 | -35.5303 | -34.798
24.4048 | -35.5307 | -34.8036
24.4109 | -35.531 | -34.8093
24.417 | -35.5314 | -34.8151
24.4231 | -35.5318 | -34.821
24.4293 | -35.5322 | -34.8267
24.4354 | -35.5326 | -34.8324
24.4415 | -35.533 | -34.8381
24.4476 | -35.5334 | -34.8438
24.4537 | -35.5338 | -34.8495
24.4598 | -35.5343 | -34.8554
24.4659 | -35.5347 | -34.8612
24.472 | -35.5351 | -34.867
24.4781 | -35.5355 | -34.8728
24.4842 | -35.5359 | -34.8785
24.4903 | -35.5362 | -34.8843
24.4964 | -35.5366 | -34.89
24.5025 | -35.537 | -34.896
24.5086 | -35.5374 | -34.9018
24.5147 | -35.5379 | -34.9077
24.5208 | -35.5383 | -34.9131
24.5269 | -35.5387 | -34.9188
24.533 | -35.5391 | -34.9243
24.5391 | -35.5395 | -34.9303
24.5452 | -35.5399 | -34.936
24.5513 | -35.5403 | -34.9417
24.5574 | -35.5408 | -34.9476
24.5635 | -35.5411 | -34.953
24.5696 | -35.5415 | -34.959
24.5758 | -35.542 | -34.9649
24.5819 | -35.5424 | -34.9708
24.588 | -35.5427 | -34.9763
24.5941 | -35.5432 | -34.9822
24.6002 | -35.5436 | -34.9878
24.6063 | -35.5439 | -34.9936
24.6124 | -35.5443 | -34.9997
24.6185 | -35.5447 | -35.0055
24.6246 | -35.5451 | -35.0112
24.6307 | -35.5455 | -35.0171
24.6368 | -35.5458 | -35.0229
24.6429 | -35.5462 | -35.0285
24.649 | -35.5466 | -35.0341
24.6551 | -35.547 | -35.0397
24.6612 | -35.5474 | -35.0453
24.6673 | -35.5478 | -35.0511
24.6734 | -35.5482 | -35.057
24.6795 | -35.5486 | -35.0628
24.6856 | -35.549 | -35.0687
24.6917 | -35.5494 | -35.0744
24.6978 | -35.5498 | -35.0799
24.7039 | -35.5502 | -35.0857
24.71 | -35.5506 | -35.0917
24.7162 | -35.551 | -35.0976
24.7223 | -35.5514 | -35.1032
24.7284 | -35.5518 | -35.1087
24.7345 | -35.5522 | -35.1144
24.7406 | -35.5526 | -35.1202
24.7467 | -35.553 | -35.1259
24.7528 | -35.5533 | -35.1316
24.7589 | -35.5537 | -35.137
24.765 | -35.5541 | -35.1422
24.7711 | -35.5545 | -35.148
24.7772 | -35.5549 | -35.1538
24.7833 | -35.5553 | -35.1595
24.7894 | -35.5557 | -35.1651
24.7955 | -35.5561 | -35.1709
24.8016 | -35.5565 | -35.1766
24.8077 | -35.5568 | -35.1824
24.8138 | -35.5573 | -35.1883
24.8199 | -35.5576 | -35.1941
24.826 | -35.5581 | -35.1996
24.8321 | -35.5585 | -35.2055
24.8382 | -35.5588 | -35.2112
24.8443 | -35.5592 | -35.2169
24.8504 | -35.5596 | -35.2224
24.8565 | -35.56 | -35.228
24.8627 | -35.5604 | -35.2339
24.8688 | -35.5608 | -35.2394
24.8749 | -35.5612 | -35.245
24.881 | -35.5616 | -35.2508
24.8871 | -35.562 | -35.2565
24.8932 | -35.5624 | -35.262
24.8993 | -35.5629 | -35.2676
24.9054 | -35.5633 | -35.2731
24.9115 | -35.5637 | -35.2789
24.9176 | -35.564 | -35.2847
24.9237 | -35.5645 | -35.2905
24.9298 | -35.5648 | -35.2962
24.9359 | -35.5653 | -35.302
24.942 | -35.5656 | -35.3075
24.9481 | -35.5661 | -35.3131
24.9542 | -35.5664 | -35.319
24.9603 | -35.5669 | -35.3244
24.9664 | -35.5673 | -35.33
24.9725 | -35.5676 | -35.3358
24.9786 | -35.568 | -35.3412
24.9847 | -35.5684 | -35.3469
24.9908 | -35.5688 | -35.3526
24.9969 | -35.5692 | -35.3584
25.0031 | -35.5696 | -35.3641
25.0092 | -35.57 | -35.3695
25.0153 | -35.5704 | -35.375
25.0214 | -35.5708 | -35.381
25.0275 | -35.5712 | -35.3866
25.0336 | -35.5716 | -35.3925
25.0397 | -35.572 | -35.3983
25.0458 | -35.5724 | -35.4038
25.0519 | -35.5728 | -35.4096
25.058 | -35.5732 | -35.4154
25.0641 | -35.5735 | -35.421
25.0702 | -35.5739 | -35.4265
25.0763 | -35.5743 | -35.4319
25.0824 | -35.5747 | -35.4373
25.0885 | -35.5751 | -35.4429
25.0946 | -35.5755 | -35.4484
25.1007 | -35.5758 | -35.4542
25.1068 | -35.5762 | -35.4598
25.1129 | -35.5766 | -35.4654
25.119 | -35.577 | -35.4708
25.1251 | -35.5774 | -35.4762
25.1312 | -35.5778 | -35.4818
25.1373 | -35.5782 | -35.4876
25.1435 | -35.5786 | -35.4936
25.1496 | -35.579 | -35.499
25.1557 | -35.5794 | -35.5045
25.1618 | -35.5798 | -35.5103
25.1679 | -35.5801 | -35.5159
25.174 | -35.5806 | -35.5215
25.1801 | -35.581 | -35.5272
25.1862 | -35.5814 | -35.5326
25.1923 | -35.5817 | -35.5381
25.1984 | -35.5822 | -35.5436
25.2045 | -35.5826 | -35.5492
25.2106 | -35.5829 | -35.5549
25.2167 | -35.5833 | -35.5607
25.2228 | -35.5836 | -35.566
25.2289 | -35.584 | -35.5717
25.235 | -35.5844 | -35.5773
25.2411 | -35.5848 | -35.5831
25.2472 | -35.5851 | -35.5886
25.2533 | -35.5855 | -35.5942
25.2594 | -35.5859 | -35.5995
25.2655 | -35.5863 | -35.6052
25.2716 | -35.5867 | -35.6108
25.2777 | -35.5871 | -35.6166
25.2838 | -35.5874 | -35.6223
25.29 | -35.5878 | -35.6277
25.2961 | -35.5882 | -35.6332
25.3022 | -35.5886 | -35.6386
25.3083 | -35.589 | -35.6442
25.3144 | -35.5894 | -35.6498
25.3205 | -35.5897 | -35.6553
25.3266 | -35.5901 | -35.6608
25.3327 | -35.5905 | -35.6662
25.3388 | -35.5909 | -35.6717
25.3449 | -35.5912 | -35.6773
25.351 | -35.5916 | -35.6829
25.3571 | -35.592 | -35.6885
25.3632 | -35.5924 | -35.6938
25.3693 | -35.5928 | -35.6992
25.3754 | -35.5932 | -35.7048
25.3815 | -35.5936 | -35.7103
25.3876 | -35.594 | -35.7161
25.3937 | -35.5944 | -35.7216
25.3998 | -35.5948 | -35.727
25.4059 | -35.5951 | -35.7328
25.412 | -35.5955 | -35.7383
25.4181 | -35.5959 | -35.7438
25.4242 | -35.5963 | -35.7495
25.4304 | -35.5967 | -35.755
25.4365 | -35.5971 | -35.7603
25.4426 | -35.5975 | -35.766
25.4487 | -35.5979 | -35.7716
25.4548 | -35.5982 | -35.7772
25.4609 | -35.5986 | -35.7828
25.467 | -35.599 | -35.7881
25.4731 | -35.5994 | -35.7936
25.4792 | -35.5998 | -35.7991
25.4853 | -35.6002 | -35.8046
25.4914 | -35.6006 | -35.8102
25.4975 | -35.601 | -35.8155
25.5036 | -35.6014 | -35.8206
25.5097 | -35.6018 | -35.826
25.5158 | -35.6022 | -35.8315
25.5219 | -35.6026 | -35.837
25.528 | -35.6029 | -35.8427
25.5341 | -35.6033 | -35.8483
25.5402 | -35.6037 | -35.8538
25.5463 | -35.6041 | -35.8594
25.5524 | -35.6045 | -35.8649
25.5585 | -35.6049 | -35.8703
25.5646 | -35.6053 | -35.8758
25.5707 | -35.6056 | -35.8812
25.5769 | -35.606 | -35.8867
25.583 | -35.6064 | -35.8923
25.5891 | -35.6068 | -35.8978
25.5952 | -35.6072 | -35.9032
25.6013 | -35.6076 | -35.9086
25.6074 | -35.608 | -35.9138
25.6135 | -35.6084 | -35.9194
25.6196 | -35.6088 | -35.9252
25.6257 | -35.6092 | -35.9308
25.6318 | -35.6096 | -35.9361
25.6379 | -35.61 | -35.9415
25.644 | -35.6103 | -35.9467
25.6501 | -35.6107 | -35.952
25.6562 | -35.6111 | -35.9575
25.6623 | -35.6115 | -35.9631
25.6684 | -35.6119 | -35.9686
25.6745 | -35.6123 | -35.9742
25.6806 | -35.6127 | -35.9794
25.6867 | -35.6131 | -35.9852
25.6928 | -35.6135 | -35.9907
25.6989 | -35.6138 | -35.9962
25.705 | -35.6142 | -36.002
25.7111 | -35.6145 | -36.0075
25.7173 | -35.6149 | -36.0127
25.7234 | -35.6153 | -36.0183
25.7295 | -35.6156 | -36.0237
25.7356 | -35.616 | -36.0293
25.7417 | -35.6164 | -36.0347
25.7478 | -35.6168 | -36.04
25.7539 | -35.6172 | -36.0453
25.76 | -35.6176 | -36.0509
25.7661 | -35.6179 | -36.0562
25.7722 | -35.6183 | -36.0618
25.7783 | -35.6187 | -36.0674
25.7844 | -35.6191 | -36.0726
25.7905 | -35.6195 | -36.0782
25.7966 | -35.6198 | -36.0834
25.8027 | -35.6202 | -36.0891
25.8088 | -35.6206 | -36.0948
25.8149 | -35.6209 | -36.1002
25.821 | -35.6213 | -36.1055
25.8271 | -35.6217 | -36.1111
25.8332 | -35.6221 | -36.1167
25.8393 | -35.6225 | -36.1223
25.8454 | -35.6229 | -36.1276
25.8515 | -35.6233 | -36.1331
25.8576 | -35.6237 | -36.1381
25.8638 | -35.624 | -36.1436
25.8699 | -35.6244 | -36.1491
25.876 | -35.6248 | -36.1548
25.8821 | -35.6252 | -36.1604
25.8882 | -35.6255 | -36.1656
25.8943 | -35.6259 | -36.1711
25.9004 | -35.6263 | -36.1766
25.9065 | -35.6267 | -36.1822
25.9126 | -35.627 | -36.1875
25.9187 | -35.6274 | -36.1928
25.9248 | -35.6278 | -36.198
25.9309 | -35.6282 | -36.2035
25.937 | -35.6286 | -36.2092
25.9431 | -35.629 | -36.2149
25.9492 | -35.6293 | -36.2202
25.9553 | -35.6297 | -36.2257
25.9614 | -35.6301 | -36.2307
25.9675 | -35.6304 | -36.2361
25.9736 | -35.6308 | -36.2413
25.9797 | -35.6312 | -36.2468
25.9858 | -35.6316 | -36.2522
25.9919 | -35.6319 | -36.2578
25.998 | -35.6323 | -36.2629
26.0042 | -35.6327 | -36.2684
26.0103 | -35.6331 | -36.2738
26.0164 | -35.6335 | -36.2791
26.0225 | -35.6339 | -36.2842
26.0286 | -35.6342 | -36.2895
26.0347 | -35.6346 | -36.2947
26.0408 | -35.635 | -36.3001
26.0469 | -35.6353 | -36.3054
26.053 | -35.6357 | -36.3109
26.0591 | -35.6361 | -36.3164
26.0652 | -35.6365 | -36.3215
26.0713 | -35.6368 | -36.327
26.0774 | -35.6372 | -36.3326
26.0835 | -35.6376 | -36.338
26.0896 | -35.6379 | -36.3433
26.0957 | -35.6383 | -36.3485
26.1018 | -35.6387 | -36.3535
26.1079 | -35.639 | -36.3585
26.114 | -35.6394 | -36.3641
26.1201 | -35.6397 | -36.3696
26.1262 | -35.64 | -36.3752
26.1323 | -35.6404 | -36.3804
26.1384 | -35.6408 | -36.3857
26.1445 | -35.6412 | -36.3912
26.1507 | -35.6416 | -36.3966
26.1568 | -35.6419 | -36.402
26.1629 | -35.6423 | -36.4071
26.169 | -35.6427 | -36.4125
26.1751 | -35.6431 | -36.4178
26.1812 | -35.6435 | -36.4231
26.1873 | -35.6438 | -36.4286
26.1934 | -35.6442 | -36.434
26.1995 | -35.6446 | -36.4395
26.2056 | -35.645 | -36.4447
26.2117 | -35.6453 | -36.4502
26.2178 | -35.6457 | -36.4555
26.2239 | -35.6461 | -36.4606
26.23 | -35.6465 | -36.4662
26.2361 | -35.6469 | -36.4714
26.2422 | -35.6472 | -36.4766
26.2483 | -35.6476 | -36.4816
26.2544 | -35.648 | -36.4869
26.2605 | -35.6484 | -36.4923
26.2666 | -35.6487 | -36.4977
26.2727 | -35.649 | -36.5029
26.2788 | -35.6494 | -36.5079
26.2849 | -35.6498 | -36.5133
26.2911 | -35.6502 | -36.5186
26.2972 | -35.6505 | -36.5241
26.3033 | -35.6509 | -36.5295
26.3094 | -35.6513 | -36.5351
26.3155 | -35.6516 | -36.5403
26.3216 | -35.652 | -36.5458
26.3277 | -35.6523 | -36.5512
26.3338 | -35.6528 | -36.5562
26.3399 | -35.6531 | -36.5614
26.346 | -35.6535 | -36.5664
26.3521 | -35.6538 | -36.5718
26.3582 | -35.6542 | -36.5771
26.3643 | -35.6546 | -36.5824
26.3704 | -35.655 | -36.5879
26.3765 | -35.6554 | -36.593
26.3826 | -35.6558 | -36.5983
26.3887 | -35.6561 | -36.6037
26.3948 | -35.6565 | -36.6091
26.4009 | -35.6568 | -36.6143
26.407 | -35.6572 | -36.6195
26.4131 | -35.6576 | -36.6246
26.4192 | -35.6579 | -36.6298
26.4253 | -35.6583 | -36.635
26.4314 | -35.6587 | -36.6401
26.4376 | -35.6591 | -36.6455
26.4437 | -35.6594 | -36.6505
26.4498 | -35.6598 | -36.656
26.4559 | -35.6602 | -36.661
26.462 | -35.6605 | -36.6663
26.4681 | -35.6609 | -36.6717
26.4742 | -35.6613 | -36.6771
26.4803 | -35.6617 | -36.6821
26.4864 | -35.662 | -36.6872
26.4925 | -35.6624 | -36.6925
26.4986 | -35.6627 | -36.6979
26.5047 | -35.6631 | -36.7034
26.5108 | -35.6635 | -36.7085
26.5169 | -35.6639 | -36.7138
26.523 | -35.6642 | -36.7191
26.5291 | -35.6646 | -36.7243
26.5352 | -35.6649 | -36.73
26.5413 | -35.6653 | -36.7355
26.5474 | -35.6657 | -36.7409
26.5535 | -35.666 | -36.7464
26.5596 | -35.6663 | -36.7518
26.5657 | -35.6667 | -36.757
26.5718 | -35.6671 | -36.7623
26.578 | -35.6674 | -36.7678
26.5841 | -35.6678 | -36.7732
26.5902 | -35.6682 | -36.7783
26.5963 | -35.6685 | -36.7834
26.6024 | -35.6689 | -36.7889
26.6085 | -35.6693 | -36.7942
26.6146 | -35.6697 | -36.7991
26.6207 | -35.67 | -36.8046
26.6268 | -35.6704 | -36.8099
26.6329 | -35.6707 | -36.8154
26.639 | -35.671 | -36.8206
26.6451 | -35.6714 | -36.826
26.6512 | -35.6718 | -36.831
26.6573 | -35.6721 | -36.8362
26.6634 | -35.6725 | -36.8412
26.6695 | -35.6728 | -36.8466
26.6756 | -35.6731 | -36.852
26.6817 | -35.6735 | -36.8575
26.6878 | -35.6738 | -36.8629
26.6939 | -35.6742 | -36.8682
26.7 | -35.6746 | -36.8731
26.7061 | -35.675 | -36.8783
26.7122 | -35.6753 | -36.8835
26.7183 | -35.6757 | -36.8887
26.7245 | -35.6761 | -36.894
26.7306 | -35.6765 | -36.8993
26.7367 | -35.6768 | -36.9044
26.7428 | -35.6772 | -36.9094
26.7489 | -35.6776 | -36.9149
26.755 | -35.6779 | -36.9201
26.7611 | -35.6783 | -36.9254
26.7672 | -35.6786 | -36.9311
26.7733 | -35.679 | -36.9364
26.7794 | -35.6793 | -36.9418
26.7855 | -35.6798 | -36.9471
26.7916 | -35.6801 | -36.9523
26.7977 | -35.6804 | -36.9575
26.8038 | -35.6808 | -36.9627
26.8099 | -35.6812 | -36.9678
26.816 | -35.6816 | -36.9731
26.8221 | -35.6819 | -36.9781
26.8282 | -35.6822 | -36.9832
26.8343 | -35.6826 | -36.9884
26.8404 | -35.683 | -36.9935
26.8465 | -35.6833 | -36.9987
26.8526 | -35.6837 | -37.004
26.8587 | -35.684 | -37.0094
26.8649 | -35.6844 | -37.0148
26.871 | -35.6848 | -37.02
26.8771 | -35.6852 | -37.0249
26.8832 | -35.6855 | -37.0301
26.8893 | -35.6859 | -37.035
26.8954 | -35.6862 | -37.0405
26.9015 | -35.6866 | -37.0458
26.9076 | -35.6869 | -37.0514
26.9137 | -35.6873 | -37.0565
26.9198 | -35.6877 | -37.0616
26.9259 | -35.688 | -37.0666
26.932 | -35.6884 | -37.0717
26.9381 | -35.6888 | -37.0773
26.9442 | -35.6891 | -37.0822
26.9503 | -35.6894 | -37.0874
26.9564 | -35.6898 | -37.0928
26.9625 | -35.6902 | -37.0979
26.9686 | -35.6906 | -37.1034
26.9747 | -35.6909 | -37.1087
26.9808 | -35.6913 | -37.1137
26.9869 | -35.6916 | -37.1189
26.993 | -35.692 | -37.124
26.9991 | -35.6923 | -37.1291
27.0052 | -35.6927 | -37.1342
27.0114 | -35.693 | -37.1395
27.0175 | -35.6933 | -37.1446
27.0236 | -35.6937 | -37.1497
27.0297 | -35.694 | -37.1549
27.0358 | -35.6944 | -37.16
27.0419 | -35.6948 | -37.1651
27.048 | -35.6952 | -37.1704
27.0541 | -35.6955 | -37.1756
27.0602 | -35.6959 | -37.1809
27.0663 | -35.6963 | -37.1864
27.0724 | -35.6966 | -37.1914
27.0785 | -35.697 | -37.1966
27.0846 | -35.6973 | -37.2016
27.0907 | -35.6977 | -37.207
27.0968 | -35.698 | -37.2123
27.1029 | -35.6984 | -37.2174
27.109 | -35.6987 | -37.2225
27.1151 | -35.6991 | -37.2276
27.1212 | -35.6994 | -37.2324
27.1273 | -35.6998 | -37.238
27.1334 | -35.7001 | -37.2435
27.1395 | -35.7005 | -37.2485
27.1456 | -35.7008 | -37.2537
27.1518 | -35.7012 | -37.2589
27.1579 | -35.7016 | -37.264
27.164 | -35.7019 | -37.269
27.1701 | -35.7023 | -37.2742
27.1762 | -35.7026 | -37.2792
27.1823 | -35.703 | -37.2845
27.1884 | -35.7033 | -37.2896
27.1945 | -35.7037 | -37.2944
27.2006 | -35.704 | -37.2996
27.2067 | -35.7044 | -37.3047
27.2128 | -35.7048 | -37.3101
27.2189 | -35.7051 | -37.3152
27.225 | -35.7054 | -37.3204
27.2311 | -35.7058 | -37.3254
27.2372 | -35.7061 | -37.3308
27.2433 | -35.7065 | -37.3359
27.2494 | -35.7069 | -37.3412
27.2555 | -35.7072 | -37.3463
27.2616 | -35.7076 | -37.3515
27.2677 | -35.7079 | -37.3565
27.2738 | -35.7083 | -37.3616
27.2799 | -35.7087 | -37.3669
27.286 | -35.709 | -37.3718
27.2921 | -35.7094 | -37.3771
27.2983 | -35.7097 | -37.3822
27.3044 | -35.71 | -37.3874
27.3105 | -35.7103 | -37.3926
27.3166 | -35.7107 | -37.3975
27.3227 | -35.7111 | -37.4028
27.3288 | -35.7114 | -37.4077
27.3349 | -35.7117 | -37.4129
27.341 | -35.7121 | -37.418
27.3471 | -35.7124 | -37.4231
27.3532 | -35.7128 | -37.4282
27.3593 | -35.7132 | -37.4336
27.3654 | -35.7135 | -37.4389
27.3715 | -35.7138 | -37.4436
27.3776 | -35.7141 | -37.4487
27.3837 | -35.7145 | -37.4535
27.3898 | -35.7149 | -37.4586
27.3959 | -35.7152 | -37.4638
27.402 | -35.7155 | -37.4686
27.4081 | -35.7159 | -37.4736
27.4142 | -35.7162 | -37.4787
27.4203 | -35.7166 | -37.4841
27.4264 | -35.7169 | -37.4891
27.4325 | -35.7173 | -37.4943
27.4387 | -35.7176 | -37.4991
27.4448 | -35.7179 | -37.5042
27.4509 | -35.7182 | -37.5093
27.457 | -35.7186 | -37.5144
27.4631 | -35.7189 | -37.5198
27.4692 | -35.7193 | -37.5248
27.4753 | -35.7196 | -37.5296
27.4814 | -35.7199 | -37.5344
27.4875 | -35.7202 | -37.5397
27.4936 | -35.7206 | -37.5449
27.4997 | -35.721 | -37.5499
27.5058 | -35.7213 | -37.5548
27.5119 | -35.7217 | -37.5599
27.518 | -35.722 | -37.565
27.5241 | -35.7224 | -37.5704
27.5302 | -35.7228 | -37.5752
27.5363 | -35.7231 | -37.5803
27.5424 | -35.7234 | -37.5853
27.5485 | -35.7238 | -37.5904
27.5546 | -35.7241 | -37.5953
27.5607 | -35.7245 | -37.6003
27.5668 | -35.7248 | -37.6053
27.5729 | -35.7251 | -37.6103
27.5791 | -35.7254 | -37.6152
27.5852 | -35.7258 | -37.6202
27.5913 | -35.7261 | -37.6254
27.5974 | -35.7265 | -37.6307
27.6035 | -35.7268 | -37.636
27.6096 | -35.7271 | -37.6408
27.6157 | -35.7275 | -37.6456
27.6218 | -35.7278 | -37.6508
27.6279 | -35.7281 | -37.656
27.634 | -35.7285 | -37.6612
27.6401 | -35.7288 | -37.6664
27.6462 | -35.7291 | -37.6717
27.6523 | -35.7294 | -37.6768
27.6584 | -35.7298 | -37.6819
27.6645 | -35.7301 | -37.6871
27.6706 | -35.7305 | -37.6921
27.6767 | -35.7308 | -37.6971
27.6828 | -35.7311 | -37.7018
27.6889 | -35.7315 | -37.7069
27.695 | -35.7318 | -37.712
27.7011 | -35.7321 | -37.717
27.7072 | -35.7325 | -37.7221
27.7133 | -35.7328 | -37.7274
27.7194 | -35.7332 | -37.7325
27.7256 | -35.7335 | -37.7375
27.7317 | -35.7339 | -37.7423
27.7378 | -35.7342 | -37.7475
27.7439 | -35.7346 | -37.7527
27.75 | -35.7349 | -37.7581
27.7561 | -35.7352 | -37.7631
27.7622 | -35.7356 | -37.7682
27.7683 | -35.7359 | -37.7734
27.7744 | -35.7363 | -37.7787
27.7805 | -35.7366 | -37.7839
27.7866 | -35.7369 | -37.7892
27.7927 | -35.7373 | -37.794
27.7988 | -35.7376 | -37.7987
27.8049 | -35.738 | -37.8039
27.811 | -35.7383 | -37.8093
27.8171 | -35.7387 | -37.8146
27.8232 | -35.7391 | -37.8194
27.8293 | -35.7394 | -37.8243
27.8354 | -35.7397 | -37.8295
27.8415 | -35.7401 | -37.8345
27.8476 | -35.7404 | -37.8394
27.8537 | -35.7407 | -37.8445
27.8598 | -35.741 | -37.8493
27.866 | -35.7414 | -37.8544
27.8721 | -35.7418 | -37.8594
27.8782 | -35.7421 | -37.8642
27.8843 | -35.7424 | -37.8693
27.8904 | -35.7428 | -37.8746
27.8965 | -35.7431 | -37.8798
27.9026 | -35.7434 | -37.8848
27.9087 | -35.7438 | -37.8901
27.9148 | -35.7441 | -37.8949
27.9209 | -35.7444 | -37.9001
27.927 | -35.7447 | -37.905
27.9331 | -35.7451 | -37.9104
27.9392 | -35.7454 | -37.9154
27.9453 | -35.7458 | -37.9201
27.9514 | -35.7461 | -37.9249
27.9575 | -35.7464 | -37.93
27.9636 | -35.7467 | -37.9351
27.9697 | -35.747 | -37.94
27.9758 | -35.7474 | -37.9451
27.9819 | -35.7477 | -37.9505
27.988 | -35.748 | -37.9555
27.9941 | -35.7483 | -37.9605
28.0002 | -35.7486 | -37.9654
28.0063 | -35.7489 | -37.9705
28.0125 | -35.7492 | -37.9755
28.0186 | -35.7495 | -37.9808
28.0247 | -35.7498 | -37.9854
28.0308 | -35.7502 | -37.9903
28.0369 | -35.7505 | -37.9951
28.043 | -35.7507 | -38.0008
28.0491 | -35.7511 | -38.006
28.0552 | -35.7514 | -38.0112
28.0613 | -35.7517 | -38.0167
28.0674 | -35.7521 | -38.0214
28.0735 | -35.7524 | -38.0264

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4914 | -41.7847
21.4908 | -35.5391 | -35.1273
16.4948 | -34.9167 | -27.7506
12.6602 | -34.5583 | -21.4648
9.7171 | -34.3351 | -16.5655
7.4582 | -34.1886 | -12.5781
5.7244 | -34.0802 | -9.35
4.3936 | -34.0118 | -6.7469
3.3722 | -33.9959 | -4.6916
2.5883 | -33.9859 | -3.1344
1.9866 | -33.979 | -1.891
1.5248 | -33.9853 | -0.89234
1.1703 | -33.9861 | -0.24012
1 | -33.995 | 0.081203
0.89822 | -33.9864 | 0.35468
0.68941 | -33.9711 | 0.93962
0.52919 | -33.9648 | 1.3915
0.40612 | -33.9707 | 1.9974
0.31171 | -33.9629 | 2.4542
0.23928 | -33.9728 | 3.1688
0.18358 | -33.9772 | 3.5854
0.14091 | -33.9935 | 4.4625
0.10819 | -34.0087 | 5.6213
0.083017 | -34.0459 | 7.0974
0.063705 | -34.0763 | 8.9721
0.048923 | -34.1519 | 11.4611
0.037479 | -34.2475 | 14.6672
0.028801 | -34.4287 | 18.693
0.022078 | -34.7416 | 24.0896
0.017023 | -35.2232 | 30.5205
0.013018 | -36.0268 | 38.1951
0.010014 | -37.1434 | 46.1018

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003774 | -1.4944
0.007549 | -27.1858
0.011323 | -35.6971
0.015098 | -41.0472
0.018872 | -44.2533
0.022647 | -46.4098
0.026421 | -48.5765
0.030196 | -50.62
0.03397 | -50.8109
0.037745 | -52.1282
0.041519 | -52.8916
0.045294 | -53.6001
0.049068 | -55.5494
0.052843 | -56.5337
0.056617 | -57.1941
0.060392 | -56.9778
0.064166 | -57.7654
0.067941 | -59.0136
0.071715 | -60.0324
0.07549 | -59.9749
0.079264 | -59.2063
0.083038 | -60.4833
0.086813 | -62.6141
0.090587 | -62.685
0.094362 | -62.2989
0.098136 | -63.2983
0.10191 | -63.2026
0.10569 | -63.4919
0.10946 | -64.2672
0.11323 | -64.2736
0.11701 | -64.6496
0.12078 | -64.6962
0.12456 | -65.3964
0.12833 | -67.0085
0.13211 | -66.8792
0.13588 | -66.9563
0.13966 | -67.236
0.14343 | -68.3448
0.14721 | -67.6744
0.15098 | -67.2369
0.15475 | -68.1851
0.15853 | -69.0316
0.1623 | -69.2752
0.16608 | -69.4904
0.16985 | -69.4895
0.17363 | -70.4731
0.1774 | -70.767
0.18118 | -70.5419
0.18495 | -70.2063
0.18872 | -70.3335
0.1925 | -70.9565
0.19627 | -71.0362
0.20005 | -70.9258
0.20382 | -71.2315
0.2076 | -72.3847
0.21137 | -72.1317
0.21515 | -72.0485
0.21892 | -71.0963
0.22269 | -72.2466
0.22647 | -72.6424
0.23024 | -72.1187
0.23402 | -72.5672
0.23779 | -72.7977
0.24157 | -72.583
0.24534 | -72.1683
0.24912 | -72.471
0.25289 | -73.4177
0.25666 | -72.911
0.26044 | -73.1894
0.26421 | -73.1271
0.26799 | -73.5455
0.27176 | -74.646
0.27554 | -75.6431
0.27931 | -74.5463
0.28309 | -74.5599
0.28686 | -75.1632
0.29063 | -74.88
0.29441 | -74.9269
0.29818 | -75.7243
0.30196 | -76.5045
0.30573 | -76.7194
0.30951 | -75.8252
0.31328 | -75.7529
0.31706 | -75.2968
0.32083 | -75.7135
0.3246 | -76.5306
0.32838 | -75.3333
0.33215 | -75.5214
0.33593 | -76.1217
0.3397 | -76.4226
0.34348 | -76.2003
0.34725 | -76.9997
0.35103 | -77.9588
0.3548 | -77.8993
0.35857 | -77.8208
0.36235 | -78.6312
0.36612 | -78.0085
0.3699 | -76.9016
0.37367 | -76.9934
0.37745 | -77.1511
0.38122 | -78.3744
0.385 | -79.0273
0.38877 | -79.104
0.39254 | -78.1874
0.39632 | -78.5899
0.40009 | -78.4159
0.40387 | -78.5684
0.40764 | -79.8682
0.41142 | -78.8851
0.41519 | -77.8189
0.41897 | -78.6899
0.42274 | -78.9522
0.42652 | -78.3027
0.43029 | -79.2077
0.43406 | -79.9855
0.43784 | -78.5974
0.44161 | -78.3762
0.44539 | -79.701
0.44916 | -80.1945
0.45294 | -79.067
0.45671 | -79.1079
0.46049 | -79.9817
0.46426 | -79.5081
0.46803 | -79.2343
0.47181 | -79.1249
0.47558 | -79.6066
0.47936 | -79.7506
0.48313 | -79.5535
0.48691 | -79.8379
0.49068 | -80.5049
0.49446 | -80.1208
0.49823 | -79.6432
0.50201 | -80.2309
0.50578 | -81.4658
0.50955 | -80.6335
0.51333 | -80.1018
0.5171 | -79.2852
0.52088 | -79.2015
0.52465 | -79.7962
0.52843 | -80.0786
0.5322 | -80.5249
0.53597 | -82.0317
0.53975 | -82.2019
0.54352 | -81.8509
0.5473 | -81.0685
0.55107 | -80.2656
0.55485 | -80.4137
0.55862 | -81.3528
0.5624 | -82.2591
0.56617 | -82.5341
0.56995 | -82.1893
0.57372 | -82.0307
0.57749 | -82.3864
0.58127 | -82.8731
0.58504 | -83.0423
0.58882 | -82.9735
0.59259 | -82.5137
0.59637 | -81.6333
0.60014 | -81.3423
0.60392 | -81.5512
0.60769 | -81.8699
0.61147 | -82.2201
0.61524 | -82.6559
0.61901 | -83.1169
0.62279 | -82.6185
0.62656 | -82.0842
0.63034 | -82.2237
0.63411 | -82.5427
0.63789 | -82.4326
0.64166 | -82.5903
0.64543 | -83.4637
0.64921 | -84.142
0.65298 | -84.4313
0.65676 | -84.3165
0.66053 | -83.9652
0.66431 | -83.8321
0.66808 | -83.861
0.67186 | -83.9659
0.67563 | -83.7179
0.67941 | -83.3617
0.68318 | -83.4398
0.68695 | -83.3534
0.69073 | -83.2657
0.6945 | -83.3818
0.69828 | -83.7231
0.70205 | -83.9897
0.70583 | -83.8165
0.7096 | -83.7432
0.71338 | -84.1792
0.71715 | -84.9136
0.72093 | -85.291
0.7247 | -84.8653
0.72847 | -83.658
0.73225 | -83.069
0.73602 | -83.6778
0.7398 | -84.3484
0.74357 | -84.5281
0.74735 | -84.4354
0.75112 | -84.0063
0.75489 | -84.1613
0.75867 | -84.3756
0.76244 | -84.4255
0.76622 | -84.5195
0.76999 | -84.4056
0.77377 | -84.2097
0.77754 | -84.084
0.78132 | -84.1946
0.78509 | -84.7231
0.78887 | -84.8422
0.79264 | -84.5624
0.79641 | -84.6842
0.80019 | -85.1028
0.80396 | -85.5193
0.80774 | -85.3506
0.81151 | -84.8566
0.81529 | -84.4746
0.81906 | -84.4691
0.82284 | -84.8071
0.82661 | -85.1666
0.83039 | -85.9715
0.83416 | -86.5307
0.83793 | -86.4325
0.84171 | -86.2668
0.84548 | -86.2775
0.84926 | -86.5
0.85303 | -86.4329
0.85681 | -85.8154
0.86058 | -85.5591
0.86435 | -85.5663
0.86813 | -85.3468
0.8719 | -84.9276
0.87568 | -84.6651
0.87945 | -84.8518
0.88323 | -84.896
0.887 | -84.9169
0.89078 | -84.8526
0.89455 | -84.9144
0.89833 | -85.2431
0.9021 | -85.9167
0.90587 | -86.3534
0.90965 | -86.3756
0.91342 | -86.4324
0.9172 | -86.142
0.92097 | -85.8716
0.92475 | -85.8958
0.92852 | -85.9235
0.93229 | -85.7112
0.93607 | -85.9736
0.93984 | -86.7919
0.94362 | -87.446
0.94739 | -87.6617
0.95117 | -87.4126
0.95494 | -87.3341
0.95872 | -87.5221
0.96249 | -87.3655
0.96627 | -87.0213
0.97004 | -86.8075
0.97381 | -86.7523
0.97759 | -86.4501
0.98136 | -86.1264
0.98514 | -85.672
0.98891 | -85.5042
0.99269 | -85.9388
0.99646 | -86.3572
1.0002 | -86.3393
1.004 | -86.1721
1.0078 | -86.364
1.0116 | -86.6749
1.0153 | -87.0598
1.0191 | -87.5146
1.0229 | -87.6371
1.0267 | -87.737
1.0304 | -87.9404
1.0342 | -88.0822
1.038 | -88.0341
1.0418 | -88.0674
1.0455 | -88.3345
1.0493 | -88.5888
1.0531 | -88.4935
1.0569 | -88.2189
1.0606 | -87.83
1.0644 | -87.5566
1.0682 | -87.3368
1.072 | -87.4123
1.0757 | -87.8377
1.0795 | -88.0185
1.0833 | -87.671
1.087 | -87.7762
1.0908 | -88.1654
1.0946 | -88.6499
1.0984 | -88.7132
1.1021 | -88.3828
1.1059 | -88.1088
1.1097 | -88.0952
1.1135 | -88.3266
1.1172 | -88.6434
1.121 | -88.8353
1.1248 | -88.8619
1.1286 | -88.6943
1.1323 | -88.3167
1.1361 | -87.9576
1.1399 | -87.8573
1.1437 | -88.0444
1.1474 | -88.3197
1.1512 | -88.4241
1.155 | -88.4181
1.1588 | -88.5629
1.1625 | -88.8658
1.1663 | -89.083
1.1701 | -89.0462
1.1739 | -88.8652
1.1776 | -88.8144
1.1814 | -89.0024
1.1852 | -89.259
1.189 | -89.2907
1.1927 | -89.1011
1.1965 | -88.9554
1.2003 | -88.8746
1.2041 | -88.6723
1.2078 | -88.3672
1.2116 | -88.1079
1.2154 | -88.0037
1.2192 | -88.0513
1.2229 | -88.1611
1.2267 | -88.3271
1.2305 | -88.5356
1.2343 | -88.7103
1.238 | -88.8844
1.2418 | -89.1072
1.2456 | -89.3001
1.2494 | -89.2932
1.2531 | -89.0788
1.2569 | -88.8926
1.2607 | -88.9779
1.2644 | -89.4485
1.2682 | -90.0999
1.272 | -90.3355
1.2758 | -89.8635
1.2795 | -89.2016
1.2833 | -88.8755
1.2871 | -88.9851
1.2909 | -89.2762
1.2946 | -89.3848
1.2984 | -89.314
1.3022 | -89.3313
1.306 | -89.4883
1.3097 | -89.5712
1.3135 | -89.4614
1.3173 | -89.2849
1.3211 | -89.1845
1.3248 | -89.1985
1.3286 | -89.2523
1.3324 | -89.2498
1.3362 | -89.2341
1.3399 | -89.3119
1.3437 | -89.5157
1.3475 | -89.754
1.3513 | -89.965
1.355 | -90.2394
1.3588 | -90.55
1.3626 | -90.6707
1.3664 | -90.5109
1.3701 | -90.1738
1.3739 | -89.8264
1.3777 | -89.6488
1.3815 | -89.7474
1.3852 | -90.0358
1.389 | -90.1532
1.3928 | -89.9665
1.3966 | -89.8095
1.4003 | -89.8431
1.4041 | -89.9822
1.4079 | -90.1524
1.4117 | -90.2175
1.4154 | -90.0888
1.4192 | -89.9436
1.423 | -89.9516
1.4268 | -90.0178
1.4305 | -90.0642
1.4343 | -90.1257
1.4381 | -90.1954
1.4418 | -90.3406
1.4456 | -90.61
1.4494 | -90.8374
1.4532 | -90.8303
1.4569 | -90.656
1.4607 | -90.5617
1.4645 | -90.6086
1.4683 | -90.6196
1.472 | -90.4987
1.4758 | -90.3654
1.4796 | -90.3604
1.4834 | -90.4947
1.4871 | -90.7331
1.4909 | -91.0775
1.4947 | -91.4283
1.4985 | -91.6078
1.5022 | -91.5789
1.506 | -91.4694
1.5098 | -91.3505
1.5136 | -91.1444
1.5173 | -90.9067
1.5211 | -90.8327
1.5249 | -91.031
1.5287 | -91.3721
1.5324 | -91.5206
1.5362 | -91.3889
1.54 | -91.2205
1.5438 | -91.171
1.5475 | -91.2343
1.5513 | -91.2937
1.5551 | -91.208
1.5589 | -90.9546
1.5626 | -90.7801
1.5664 | -90.9059
1.5702 | -91.1739
1.574 | -91.2425
1.5777 | -91.0727
1.5815 | -90.8978
1.5853 | -90.8798
1.5891 | -90.9572
1.5928 | -91.0559
1.5966 | -91.1558
1.6004 | -91.2893
1.6042 | -91.5229
1.6079 | -91.7342
1.6117 | -91.7134
1.6155 | -91.5209
1.6193 | -91.3633
1.623 | -91.3115
1.6268 | -91.2317
1.6306 | -91.0283
1.6343 | -90.9546
1.6381 | -91.0524
1.6419 | -91.3071
1.6457 | -91.5505
1.6494 | -91.6129
1.6532 | -91.5583
1.657 | -91.5637
1.6608 | -91.7106
1.6645 | -91.9343
1.6683 | -92.0741
1.6721 | -92.0203
1.6759 | -91.8694
1.6796 | -91.7747
1.6834 | -91.7909
1.6872 | -91.8643
1.691 | -91.9751
1.6947 | -92.1066
1.6985 | -92.2572
1.7023 | -92.4501
1.7061 | -92.6708
1.7098 | -92.7902
1.7136 | -92.7619
1.7174 | -92.6329
1.7212 | -92.4513
1.7249 | -92.1798
1.7287 | -91.8581
1.7325 | -91.5751
1.7363 | -91.3783
1.74 | -91.2783
1.7438 | -91.2846
1.7476 | -91.3933
1.7514 | -91.5475
1.7551 | -91.6665
1.7589 | -91.7794
1.7627 | -91.9458
1.7665 | -92.1346
1.7702 | -92.2293
1.774 | -92.1691
1.7778 | -92.025
1.7816 | -91.9075
1.7853 | -91.9107
1.7891 | -92.0404
1.7929 | -92.1686
1.7967 | -92.1482
1.8004 | -92.0044
1.8042 | -91.9085
1.808 | -91.9481
1.8117 | -92.0811
1.8155 | -92.2323
1.8193 | -92.3513
1.8231 | -92.4202
1.8268 | -92.4363
1.8306 | -92.4041
1.8344 | -92.3201
1.8382 | -92.1874
1.8419 | -92.0873
1.8457 | -92.0932
1.8495 | -92.2184
1.8533 | -92.416
1.857 | -92.635
1.8608 | -92.8445
1.8646 | -92.9791
1.8684 | -92.9879
1.8721 | -92.873
1.8759 | -92.6726
1.8797 | -92.4566
1.8835 | -92.3152
1.8872 | -92.284
1.891 | -92.3404
1.8948 | -92.4596
1.8986 | -92.6706
1.9023 | -92.9561
1.9061 | -93.2146
1.9099 | -93.3211
1.9137 | -93.28
1.9174 | -93.1869
1.9212 | -93.1398
1.925 | -93.1819
1.9288 | -93.2907
1.9325 | -93.3865
1.9363 | -93.4008
1.9401 | -93.3129
1.9439 | -93.1345
1.9476 | -92.8989
1.9514 | -92.6659
1.9552 | -92.5323
1.959 | -92.5837
1.9627 | -92.8531
1.9665 | -93.2463
1.9703 | -93.6373
1.9741 | -93.9213
1.9778 | -94.0464
1.9816 | -93.9853
1.9854 | -93.832
1.9891 | -93.7256
1.9929 | -93.7299
1.9967 | -93.8172
2.0005 | -93.9381
2.0042 | -94.0348
2.008 | -94.0835
2.0118 | -94.035
2.0156 | -93.865
2.0193 | -93.6286
2.0231 | -93.4678
2.0269 | -93.4836
2.0307 | -93.6272
2.0344 | -93.7523
2.0382 | -93.8021
2.042 | -93.8382
2.0458 | -93.8779
2.0495 | -93.869
2.0533 | -93.8133
2.0571 | -93.767
2.0609 | -93.7612
2.0646 | -93.7919
2.0684 | -93.8432
2.0722 | -93.9097
2.076 | -93.9949
2.0797 | -94.0486
2.0835 | -93.9975
2.0873 | -93.8455
2.0911 | -93.7045
2.0948 | -93.6312
2.0986 | -93.5695
2.1024 | -93.4275
2.1062 | -93.1711
2.1099 | -92.9297
2.1137 | -92.8978
2.1175 | -93.1281
2.1213 | -93.4761
2.125 | -93.7027
2.1288 | -93.7567
2.1326 | -93.7549
2.1364 | -93.7772
2.1401 | -93.8254
2.1439 | -93.845
2.1477 | -93.8357
2.1515 | -93.887
2.1552 | -94.0569
2.159 | -94.2898
2.1628 | -94.4551
2.1665 | -94.5254
2.1703 | -94.5887
2.1741 | -94.6903
2.1779 | -94.6843
2.1816 | -94.6253
2.1854 | -94.4984
2.1892 | -94.3491
2.193 | -94.2251
2.1967 | -94.1482
2.2005 | -94.1112
2.2043 | -94.0831
2.2081 | -94.0273
2.2118 | -93.9407
2.2156 | -93.8541
2.2194 | -93.8064
2.2232 | -93.8147
2.2269 | -93.8624
2.2307 | -93.9395
2.2345 | -94.0609
2.2383 | -94.2294
2.242 | -94.4346
2.2458 | -94.6451
2.2496 | -94.8154
2.2534 | -94.9245
2.2571 | -94.9813
2.2609 | -95.0052
2.2647 | -95.003
2.2685 | -94.953
2.2722 | -94.8287
2.276 | -94.6647
2.2798 | -94.545
2.2836 | -94.5102
2.2873 | -94.5626
2.2911 | -94.6471
2.2949 | -94.6947
2.2987 | -94.6876
2.3024 | -94.6486
2.3062 | -94.6142
2.31 | -94.6121
2.3138 | -94.6318
2.3175 | -94.6372
2.3213 | -94.5911
2.3251 | -94.4592
2.3289 | -94.2516
2.3326 | -94.0404
2.3364 | -93.9211
2.3402 | -93.9359
2.3439 | -94.0582
2.3477 | -94.2195
2.3515 | -94.3643
2.3553 | -94.4749
2.359 | -94.5653
2.3628 | -94.6463
2.3666 | -94.7408
2.3704 | -94.8126
2.3741 | -94.8106
2.3779 | -94.7157
2.3817 | -94.5667
2.3855 | -94.4372
2.3892 | -94.3644
2.393 | -94.3324
2.3968 | -94.335
2.4006 | -94.357
2.4043 | -94.385
2.4081 | -94.4079
2.4119 | -94.4264
2.4157 | -94.4209
2.4194 | -94.3819
2.4232 | -94.3329
2.427 | -94.2977
2.4308 | -94.28
2.4345 | -94.2725
2.4383 | -94.2836
2.4421 | -94.334
2.4459 | -94.453
2.4496 | -94.6516
2.4534 | -94.9198
2.4572 | -95.2059
2.461 | -95.4108
2.4647 | -95.4575
2.4685 | -95.3341
2.4723 | -95.1295
2.4761 | -94.9523
2.4798 | -94.8576
2.4836 | -94.8675
2.4874 | -94.9632
2.4912 | -95.1059
2.4949 | -95.2183
2.4987 | -95.2676
2.5025 | -95.2552
2.5063 | -95.2035
2.51 | -95.1499
2.5138 | -95.133
2.5176 | -95.181
2.5213 | -95.2919
2.5251 | -95.3958
2.5289 | -95.4144
2.5327 | -95.3225
2.5364 | -95.1669
2.5402 | -94.9774
2.544 | -94.8171
2.5478 | -94.7356
2.5515 | -94.7327
2.5553 | -94.7625
2.5591 | -94.798
2.5629 | -94.8405
2.5666 | -94.921
2.5704 | -95.0603
2.5742 | -95.2768
2.578 | -95.5608
2.5817 | -95.8722
2.5855 | -96.1364
2.5893 | -96.3118
2.5931 | -96.4004
2.5968 | -96.4182
2.6006 | -96.3576
2.6044 | -96.2238
2.6082 | -96.0272
2.6119 | -95.7758
2.6157 | -95.5147
2.6195 | -95.2952
2.6233 | -95.148
2.627 | -95.0897
2.6308 | -95.1336
2.6346 | -95.298
2.6384 | -95.5552
2.6421 | -95.846
2.6459 | -96.1001
2.6497 | -96.2742
2.6535 | -96.3509
2.6572 | -96.3124
2.661 | -96.218
2.6648 | -96.1571
2.6686 | -96.1686
2.6723 | -96.2255
2.6761 | -96.2542
2.6799 | -96.2326
2.6837 | -96.1564
2.6874 | -96.0435
2.6912 | -95.9179
2.695 | -95.7856
2.6987 | -95.6483
2.7025 | -95.5268
2.7063 | -95.4409
2.7101 | -95.4107
2.7138 | -95.4369
2.7176 | -95.5051
2.7214 | -95.5837
2.7252 | -95.6537
2.7289 | -95.7311
2.7327 | -95.8162
2.7365 | -95.8972
2.7403 | -95.9808
2.744 | -96.0546
2.7478 | -96.0914
2.7516 | -96.0941
2.7554 | -96.0546
2.7591 | -95.9784
2.7629 | -95.8923
2.7667 | -95.8429
2.7705 | -95.8446
2.7742 | -95.8949
2.778 | -95.9473
2.7818 | -95.9607
2.7856 | -95.9302
2.7893 | -95.8732
2.7931 | -95.827
2.7969 | -95.8087
2.8007 | -95.8098
2.8044 | -95.8318
2.8082 | -95.8668
2.812 | -95.9181
2.8158 | -95.979
2.8195 | -96.0433
2.8233 | -96.0948
2.8271 | -96.1488
2.8309 | -96.2353
2.8346 | -96.3711
2.8384 | -96.5466
2.8422 | -96.7186
2.846 | -96.8454
2.8497 | -96.8992
2.8535 | -96.8944
2.8573 | -96.8734
2.8611 | -96.8515
2.8648 | -96.8346
2.8686 | -96.8427
2.8724 | -96.8543
2.8761 | -96.8506
2.8799 | -96.8138
2.8837 | -96.7461
2.8875 | -96.6676
2.8912 | -96.6079
2.895 | -96.5771
2.8988 | -96.5793
2.9026 | -96.5775
2.9063 | -96.5506
2.9101 | -96.4887
2.9139 | -96.3996
2.9177 | -96.2928
2.9214 | -96.1846
2.9252 | -96.0669
2.929 | -95.9369
2.9328 | -95.8131
2.9365 | -95.7323
2.9403 | -95.7273
2.9441 | -95.8073
2.9479 | -95.9501
2.9516 | -96.1096
2.9554 | -96.2456
2.9592 | -96.3252
2.963 | -96.3494
2.9667 | -96.3356
2.9705 | -96.2971
2.9743 | -96.2313
2.9781 | -96.1553
2.9818 | -96.102
2.9856 | -96.0858
2.9894 | -96.1124
2.9932 | -96.1863
2.9969 | -96.304
3.0007 | -96.4468
3.0045 | -96.5869
3.0083 | -96.717
3.012 | -96.8103
3.0158 | -96.8409
3.0196 | -96.8434
3.0234 | -96.8279
3.0271 | -96.7944
3.0309 | -96.7371
3.0347 | -96.6742
3.0385 | -96.6081
3.0422 | -96.5215
3.046 | -96.4141
3.0498 | -96.3189
3.0535 | -96.2608
3.0573 | -96.2681
3.0611 | -96.3675
3.0649 | -96.5514
3.0686 | -96.7747
3.0724 | -97.0269
3.0762 | -97.2745
3.08 | -97.4908
3.0837 | -97.6237
3.0875 | -97.6593
3.0913 | -97.5974
3.0951 | -97.4374
3.0988 | -97.2035
3.1026 | -96.9602
3.1064 | -96.742
3.1102 | -96.552
3.1139 | -96.4094
3.1177 | -96.3132
3.1215 | -96.2718
3.1253 | -96.3013
3.129 | -96.4012
3.1328 | -96.5627
3.1366 | -96.7337
3.1404 | -96.8808
3.1441 | -96.9858
3.1479 | -97.022
3.1517 | -96.9913
3.1555 | -96.9257
3.1592 | -96.8743
3.163 | -96.8722
3.1668 | -96.9382
3.1706 | -97.0412
3.1743 | -97.1231
3.1781 | -97.1376
3.1819 | -97.0485
3.1857 | -96.8801
3.1894 | -96.6434
3.1932 | -96.3618
3.197 | -96.0879
3.2008 | -95.8792
3.2045 | -95.7698
3.2083 | -95.7667
3.2121 | -95.8429
3.2159 | -95.9758
3.2196 | -96.1632
3.2234 | -96.3902
3.2272 | -96.6384
3.231 | -96.8896
3.2347 | -97.1154
3.2385 | -97.3056
3.2423 | -97.464
3.246 | -97.5836
3.2498 | -97.6395
3.2536 | -97.6623
3.2574 | -97.6603
3.2611 | -97.6584
3.2649 | -97.6293
3.2687 | -97.6082
3.2725 | -97.5669
3.2762 | -97.5212
3.28 | -97.4838
3.2838 | -97.4534
3.2876 | -97.4255
3.2913 | -97.4018
3.2951 | -97.3578
3.2989 | -97.3
3.3027 | -97.2422
3.3064 | -97.1884
3.3102 | -97.161
3.314 | -97.1578
3.3178 | -97.1788
3.3215 | -97.215
3.3253 | -97.2439
3.3291 | -97.2634
3.3329 | -97.2579
3.3366 | -97.2208
3.3404 | -97.1685
3.3442 | -97.1052
3.348 | -97.0414
3.3517 | -97.0025
3.3555 | -96.9842
3.3593 | -96.9926
3.3631 | -97.0317
3.3668 | -97.0919
3.3706 | -97.1725
3.3744 | -97.2714
3.3782 | -97.3919
3.3819 | -97.5278
3.3857 | -97.6631
3.3895 | -97.7868
3.3933 | -97.8786
3.397 | -97.9329
3.4008 | -97.944
3.4046 | -97.9177
3.4084 | -97.846
3.4121 | -97.7496
3.4159 | -97.6348
3.4197 | -97.5132
3.4234 | -97.4034
3.4272 | -97.3214
3.431 | -97.2763
3.4348 | -97.2694
3.4385 | -97.294
3.4423 | -97.3435
3.4461 | -97.4121
3.4499 | -97.4982
3.4536 | -97.5911
3.4574 | -97.6805
3.4612 | -97.7489
3.465 | -97.7884
3.4687 | -97.7971
3.4725 | -97.7819
3.4763 | -97.7508
3.4801 | -97.7156
3.4838 | -97.6845
3.4876 | -97.6565
3.4914 | -97.6365
3.4952 | -97.6363
3.4989 | -97.6445
3.5027 | -97.6626
3.5065 | -97.6934
3.5103 | -97.7345
3.514 | -97.7917
3.5178 | -97.8489
3.5216 | -97.8979
3.5254 | -97.9332
3.5291 | -97.9666
3.5329 | -97.9899
3.5367 | -98.0128
3.5405 | -98.0436
3.5442 | -98.083
3.548 | -98.1114
3.5518 | -98.1123
3.5556 | -98.077
3.5593 | -98.0028
3.5631 | -97.8987
3.5669 | -97.7986
3.5707 | -97.7257
3.5744 | -97.7068
3.5782 | -97.7432
3.582 | -97.8397
3.5858 | -97.9791
3.5895 | -98.1529
3.5933 | -98.3226
3.5971 | -98.4546
3.6008 | -98.5314
3.6046 | -98.5529
3.6084 | -98.5277
3.6122 | -98.4701
3.6159 | -98.4104
3.6197 | -98.3505
3.6235 | -98.2981
3.6273 | -98.2374
3.631 | -98.1783
3.6348 | -98.1312
3.6386 | -98.094
3.6424 | -98.0869
3.6461 | -98.1193
3.6499 | -98.1812
3.6537 | -98.2478
3.6575 | -98.2889
3.6612 | -98.2776
3.665 | -98.2242
3.6688 | -98.1523
3.6726 | -98.0736
3.6763 | -98.0113
3.6801 | -97.9657
3.6839 | -97.9544
3.6877 | -97.9786
3.6914 | -98.0584
3.6952 | -98.1974
3.699 | -98.3716
3.7028 | -98.5385
3.7065 | -98.6714
3.7103 | -98.7333
3.7141 | -98.7021
3.7179 | -98.6006
3.7216 | -98.4502
3.7254 | -98.2878
3.7292 | -98.1579
3.733 | -98.0938
3.7367 | -98.1097
3.7405 | -98.1874
3.7443 | -98.2999
3.7481 | -98.4141
3.7518 | -98.4913
3.7556 | -98.5252
3.7594 | -98.5189
3.7632 | -98.4886
3.7669 | -98.469
3.7707 | -98.4748
3.7745 | -98.4991
3.7782 | -98.5462
3.782 | -98.6215
3.7858 | -98.7178
3.7896 | -98.8238
3.7933 | -98.9199
3.7971 | -98.9933
3.8009 | -99.0157
3.8047 | -98.9762
3.8084 | -98.8965
3.8122 | -98.8474
3.816 | -98.8022
3.8198 | -98.7713
3.8235 | -98.7549
3.8273 | -98.7396
3.8311 | -98.7317
3.8349 | -98.7405
3.8386 | -98.7465
3.8424 | -98.7516
3.8462 | -98.7399
3.85 | -98.7097
3.8537 | -98.6683
3.8575 | -98.6196
3.8613 | -98.5671
3.8651 | -98.5159
3.8688 | -98.4649
3.8726 | -98.4224
3.8764 | -98.3924
3.8802 | -98.3808
3.8839 | -98.389
3.8877 | -98.4247
3.8915 | -98.4866
3.8953 | -98.5602
3.899 | -98.6378
3.9028 | -98.7235
3.9066 | -98.8107
3.9104 | -98.9024
3.9141 | -98.9909
3.9179 | -99.0612
3.9217 | -99.1125
3.9255 | -99.1373
3.9292 | -99.1267
3.933 | -99.0716
3.9368 | -98.9888
3.9406 | -98.901
3.9443 | -98.8372
3.9481 | -98.8144
3.9519 | -98.8286
3.9556 | -98.8787
3.9594 | -98.9452
3.9632 | -99.0104
3.967 | -99.0623
3.9707 | -99.0798
3.9745 | -99.0651
3.9783 | -99.0164
3.9821 | -98.9424
3.9858 | -98.8513
3.9896 | -98.7514
3.9934 | -98.6551
3.9972 | -98.5786
4.0009 | -98.5225
4.0047 | -98.4882
4.0085 | -98.4769
4.0123 | -98.4892
4.016 | -98.5292
4.0198 | -98.5947
4.0236 | -98.6969
4.0274 | -98.8255
4.0311 | -98.97
4.0349 | -99.1208
4.0387 | -99.2555
4.0425 | -99.3644
4.0462 | -99.4291
4.05 | -99.4648
4.0538 | -99.483
4.0576 | -99.4844
4.0613 | -99.4725
4.0651 | -99.4668
4.0689 | -99.4694
4.0727 | -99.4854
4.0764 | -99.526
4.0802 | -99.5841
4.084 | -99.6506
4.0878 | -99.7088
4.0915 | -99.7585
4.0953 | -99.7972
4.0991 | -99.828
4.1029 | -99.8512
4.1066 | -99.855
4.1104 | -99.8467
4.1142 | -99.8081
4.118 | -99.7549
4.1217 | -99.6865
4.1255 | -99.6155
4.1293 | -99.5646
4.133 | -99.5364
4.1368 | -99.5245
4.1406 | -99.5214
4.1444 | -99.525
4.1481 | -99.5315
4.1519 | -99.5404
4.1557 | -99.548
4.1595 | -99.5602
4.1632 | -99.5726
4.167 | -99.5695
4.1708 | -99.5433
4.1746 | -99.4815
4.1783 | -99.3953
4.1821 | -99.2954
4.1859 | -99.2031
4.1897 | -99.1258
4.1934 | -99.0582
4.1972 | -99.011
4.201 | -98.9869
4.2048 | -98.9834
4.2085 | -98.9962
4.2123 | -99.0222
4.2161 | -99.0544
4.2199 | -99.0854
4.2236 | -99.0987
4.2274 | -99.0906
4.2312 | -99.0652
4.235 | -99.0283
4.2387 | -98.9839
4.2425 | -98.9372
4.2463 | -98.9061
4.2501 | -98.8932
4.2538 | -98.8983
4.2576 | -98.9271
4.2614 | -98.9756
4.2652 | -99.0417
4.2689 | -99.1163
4.2727 | -99.207
4.2765 | -99.3064
4.2803 | -99.4096
4.284 | -99.5173
4.2878 | -99.6197
4.2916 | -99.7011
4.2954 | -99.7635
4.2991 | -99.7829
4.3029 | -99.7601
4.3067 | -99.7031
4.3105 | -99.6271
4.3142 | -99.5536
4.318 | -99.4908
4.3218 | -99.4477
4.3255 | -99.4285
4.3293 | -99.4314
4.3331 | -99.454
4.3369 | -99.4799
4.3406 | -99.508
4.3444 | -99.5339
4.3482 | -99.5523
4.352 | -99.5584
4.3557 | -99.5475
4.3595 | -99.5468
4.3633 | -99.5324
4.3671 | -99.5117
4.3708 | -99.49
4.3746 | -99.4667
4.3784 | -99.4416
4.3822 | -99.4257
4.3859 | -99.4249
4.3897 | -99.4331
4.3935 | -99.4479
4.3973 | -99.4622
4.401 | -99.4863
4.4048 | -99.5174
4.4086 | -99.555
4.4124 | -99.5979
4.4161 | -99.6487
4.4199 | -99.7082
4.4237 | -99.7552
4.4275 | -99.7863
4.4312 | -99.8049
4.435 | -99.807
4.4388 | -99.7898
4.4426 | -99.7409
4.4463 | -99.6785
4.4501 | -99.6167
4.4539 | -99.5672
4.4577 | -99.5301
4.4614 | -99.5136
4.4652 | -99.514
4.469 | -99.5322
4.4728 | -99.5532
4.4765 | -99.5799
4.4803 | -99.6116
4.4841 | -99.6517
4.4879 | -99.6962
4.4916 | -99.7531
4.4954 | -99.819
4.4992 | -99.8777
4.5029 | -99.9329
4.5067 | -99.9819
4.5105 | -100.0249
4.5143 | -100.0589
4.518 | -100.0789
4.5218 | -100.0846
4.5256 | -100.0762
4.5294 | -100.0647
4.5331 | -100.0519
4.5369 | -100.0451
4.5407 | -100.0435
4.5445 | -100.0554
4.5482 | -100.0749
4.552 | -100.1024
4.5558 | -100.1245
4.5596 | -100.1456
4.5633 | -100.1586
4.5671 | -100.1648
4.5709 | -100.1694
4.5747 | -100.1675
4.5784 | -100.1562
4.5822 | -100.1448
4.586 | -100.1224
4.5898 | -100.083
4.5935 | -100.0253
4.5973 | -99.9494
4.6011 | -99.8635
4.6049 | -99.7701
4.6086 | -99.6697
4.6124 | -99.577
4.6162 | -99.5057
4.62 | -99.4568
4.6237 | -99.4357
4.6275 | -99.4536
4.6313 | -99.4988
4.6351 | -99.571
4.6388 | -99.6664
4.6426 | -99.7718
4.6464 | -99.8806
4.6502 | -99.9889
4.6539 | -100.0743
4.6577 | -100.1389
4.6615 | -100.1676
4.6653 | -100.1743
4.669 | -100.156
4.6728 | -100.1134
4.6766 | -100.0694
4.6803 | -100.0361
4.6841 | -99.9987
4.6879 | -99.9665
4.6917 | -99.9515
4.6954 | -99.9428
4.6992 | -99.9387
4.703 | -99.9462
4.7068 | -99.9478
4.7105 | -99.9488
4.7143 | -99.9499
4.7181 | -99.9531
4.7219 | -99.9703
4.7256 | -100.0121
4.7294 | -100.0781
4.7332 | -100.1679
4.737 | -100.2702
4.7407 | -100.3798
4.7445 | -100.4843
4.7483 | -100.5701
4.7521 | -100.6403
4.7558 | -100.6974
4.7596 | -100.7212
4.7634 | -100.7156
4.7672 | -100.695
4.7709 | -100.6537
4.7747 | -100.5879
4.7785 | -100.5106
4.7823 | -100.4186
4.786 | -100.3203
4.7898 | -100.2313
4.7936 | -100.1459
4.7974 | -100.0742
4.8011 | -100.0163
4.8049 | -99.9734
4.8087 | -99.9438
4.8125 | -99.9374
4.8162 | -99.9501
4.82 | -99.9846
4.8238 | -100.034
4.8276 | -100.095
4.8313 | -100.1636
4.8351 | -100.221
4.8389 | -100.2571
4.8427 | -100.2753
4.8464 | -100.2766
4.8502 | -100.2621
4.854 | -100.2423
4.8577 | -100.2153
4.8615 | -100.1734
4.8653 | -100.1215
4.8691 | -100.0574
4.8728 | -100.0009
4.8766 | -99.9514
4.8804 | -99.9081
4.8842 | -99.8717
4.8879 | -99.8503
4.8917 | -99.8352
4.8955 | -99.8324
4.8993 | -99.864
4.903 | -99.882
4.9068 | -99.9113
4.9106 | -99.9549
4.9144 | -100.0096
4.9181 | -100.0638
4.9219 | -100.12
4.9257 | -100.1733
4.9295 | -100.2127
4.9332 | -100.2385
4.937 | -100.2593
4.9408 | -100.2717
4.9446 | -100.2795
4.9483 | -100.2813
4.9521 | -100.2893
4.9559 | -100.3137
4.9597 | -100.3452
4.9634 | -100.369
4.9672 | -100.3974
4.971 | -100.4284
4.9748 | -100.4503
4.9785 | -100.4572
4.9823 | -100.44
4.9861 | -100.3939
4.9899 | -100.3322
4.9936 | -100.2615
4.9974 | -100.1967
5.0012 | -100.1291
5.005 | -100.073
5.0087 | -100.0359
5.0125 | -100.0222
5.0163 | -100.029
5.0201 | -100.0669
5.0238 | -100.1276
5.0276 | -100.2036
5.0314 | -100.2943
5.0351 | -100.3874
5.0389 | -100.4904
5.0427 | -100.6003
5.0465 | -100.7008
5.0502 | -100.7954
5.054 | -100.8861
5.0578 | -100.978
5.0616 | -101.0741
5.0653 | -101.1561
5.0691 | -101.2212
5.0729 | -101.2809
5.0767 | -101.33
5.0804 | -101.3713
5.0842 | -101.392
5.088 | -101.3908
5.0918 | -101.3691
5.0955 | -101.322
5.0993 | -101.2512
5.1031 | -101.1742
5.1069 | -101.1045
5.1106 | -101.0415
5.1144 | -100.9905
5.1182 | -100.9505
5.122 | -100.9282
5.1257 | -100.9184
5.1295 | -100.9236
5.1333 | -100.9406
5.1371 | -100.9608
5.1408 | -100.9781
5.1446 | -101.0045
5.1484 | -101.0262
5.1522 | -101.0295
5.1559 | -101.0205
5.1597 | -100.9986
5.1635 | -100.9724
5.1673 | -100.939
5.171 | -100.8985
5.1748 | -100.8601
5.1786 | -100.8209
5.1824 | -100.7799
5.1861 | -100.7415
5.1899 | -100.6954
5.1937 | -100.6545
5.1975 | -100.6241
5.2012 | -100.6036
5.205 | -100.5968
5.2088 | -100.6019
5.2126 | -100.6288
5.2163 | -100.6688
5.2201 | -100.7158
5.2239 | -100.7638
5.2276 | -100.8121
5.2314 | -100.8673
5.2352 | -100.9296
5.239 | -100.9914
5.2427 | -101.0489
5.2465 | -101.1047
5.2503 | -101.161
5.2541 | -101.2082
5.2578 | -101.2528
5.2616 | -101.2821
5.2654 | -101.304
5.2692 | -101.3338
5.2729 | -101.3471
5.2767 | -101.3425
5.2805 | -101.3255
5.2843 | -101.3111
5.288 | -101.2935
5.2918 | -101.2727
5.2956 | -101.2542
5.2994 | -101.24
5.3031 | -101.2251
5.3069 | -101.2122
5.3107 | -101.1977
5.3145 | -101.1899
5.3182 | -101.1998
5.322 | -101.2161
5.3258 | -101.2461
5.3296 | -101.2898
5.3333 | -101.3369
5.3371 | -101.3751
5.3409 | -101.415
5.3447 | -101.4415
5.3484 | -101.447
5.3522 | -101.4443
5.356 | -101.4274
5.3598 | -101.4012
5.3635 | -101.3672
5.3673 | -101.3423
5.3711 | -101.3366
5.3749 | -101.3538
5.3786 | -101.3903
5.3824 | -101.4491
5.3862 | -101.5168
5.3899 | -101.5832
5.3937 | -101.6381
5.3975 | -101.6819
5.4013 | -101.7074
5.405 | -101.7031
5.4088 | -101.6745
5.4126 | -101.6277
5.4164 | -101.5649
5.4201 | -101.4891
5.4239 | -101.401
5.4277 | -101.3223
5.4315 | -101.2527
5.4352 | -101.2038
5.439 | -101.1725
5.4428 | -101.1565
5.4466 | -101.1314
5.4503 | -101.1114
5.4541 | -101.0952
5.4579 | -101.0872
5.4617 | -101.0849
5.4654 | -101.0827
5.4692 | -101.0871
5.473 | -101.0905
5.4768 | -101.0849
5.4805 | -101.0775
5.4843 | -101.0704
5.4881 | -101.0577
5.4919 | -101.0408
5.4956 | -101.0258
5.4994 | -101.0142
5.5032 | -101.0021
5.507 | -101.0023
5.5107 | -101.0038
5.5145 | -101.0162
5.5183 | -101.0349
5.5221 | -101.0595
5.5258 | -101.1019
5.5296 | -101.1481
5.5334 | -101.1931
5.5372 | -101.2288
5.5409 | -101.2505
5.5447 | -101.2625
5.5485 | -101.2662
5.5523 | -101.2603
5.556 | -101.2401
5.5598 | -101.2197
5.5636 | -101.1984
5.5674 | -101.1744
5.5711 | -101.1615
5.5749 | -101.1653
5.5787 | -101.1777
5.5824 | -101.2071
5.5862 | -101.2503
5.59 | -101.3022
5.5938 | -101.3636
5.5975 | -101.4297
5.6013 | -101.4954
5.6051 | -101.5496
5.6089 | -101.6001
5.6126 | -101.6322
5.6164 | -101.6516
5.6202 | -101.6533
5.624 | -101.6378
5.6277 | -101.6035
5.6315 | -101.5468
5.6353 | -101.4742
5.6391 | -101.3944
5.6428 | -101.3192
5.6466 | -101.2433
5.6504 | -101.1776
5.6542 | -101.1212
5.6579 | -101.0731
5.6617 | -101.0378
5.6655 | -101.0183
5.6693 | -101.0097
5.673 | -101.0128
5.6768 | -101.0333
5.6806 | -101.0658
5.6844 | -101.108
5.6881 | -101.1601
5.6919 | -101.2276
5.6957 | -101.2927
5.6995 | -101.3622
5.7032 | -101.4355
5.707 | -101.5125
5.7108 | -101.5955
5.7146 | -101.6762
5.7183 | -101.7459
5.7221 | -101.8218
5.7259 | -101.8901
5.7297 | -101.944
5.7334 | -101.9858
5.7372 | -102.013
5.741 | -102.0188
5.7448 | -102.0134
5.7485 | -101.9998
5.7523 | -101.9781
5.7561 | -101.9586
5.7598 | -101.9436
5.7636 | -101.9351
5.7674 | -101.9346
5.7712 | -101.9493
5.7749 | -101.9649
5.7787 | -101.989
5.7825 | -102.0171
5.7863 | -102.0376
5.79 | -102.0435
5.7938 | -102.0294
5.7976 | -101.9943
5.8014 | -101.9434
5.8051 | -101.8871
5.8089 | -101.8074
5.8127 | -101.7199
5.8165 | -101.634
5.8202 | -101.5533
5.824 | -101.4823
5.8278 | -101.4285
5.8316 | -101.3908
5.8353 | -101.3716
5.8391 | -101.3729
5.8429 | -101.3916
5.8467 | -101.4301
5.8504 | -101.4801
5.8542 | -101.5402
5.858 | -101.6083
5.8618 | -101.6829
5.8655 | -101.7506
5.8693 | -101.8142
5.8731 | -101.8736
5.8769 | -101.924
5.8806 | -101.9693
5.8844 | -101.9969
5.8882 | -102.0092
5.892 | -102.001
5.8957 | -101.9835
5.8995 | -101.958
5.9033 | -101.9234
5.9071 | -101.8907
5.9108 | -101.8583
5.9146 | -101.8358
5.9184 | -101.8279
5.9222 | -101.8402
5.9259 | -101.8721
5.9297 | -101.9208
5.9335 | -101.9892
5.9372 | -102.0674
5.941 | -102.1346
5.9448 | -102.1817
5.9486 | -102.2132
5.9523 | -102.221
5.9561 | -102.2023
5.9599 | -102.1607
5.9637 | -102.1096
5.9674 | -102.0484
5.9712 | -101.9958
5.975 | -101.9554
5.9788 | -101.9355
5.9825 | -101.9393
5.9863 | -102.0037
5.9901 | -102.0314
5.9939 | -102.0694
5.9976 | -102.1084
6.0014 | -102.1462
6.0052 | -102.175
6.009 | -102.1923
6.0127 | -102.194
6.0165 | -102.1757
6.0203 | -102.1422
6.0241 | -102.096
6.0278 | -102.0388
6.0316 | -101.9925
6.0354 | -101.9515
6.0392 | -101.9098
6.0429 | -101.8807
6.0467 | -101.8643
6.0505 | -101.8553
6.0543 | -101.8513
6.058 | -101.8571
6.0618 | -101.8731
6.0656 | -101.8957
6.0694 | -101.9192
6.0731 | -101.9491
6.0769 | -101.9854
6.0807 | -102.0276
6.0845 | -102.0718
6.0882 | -102.1158
6.092 | -102.1665
6.0958 | -102.2129
6.0996 | -102.2524
6.1033 | -102.2874
6.1071 | -102.3129
6.1109 | -102.3271
6.1147 | -102.3238
6.1184 | -102.3019
6.1222 | -102.2824
6.126 | -102.2473
6.1297 | -102.1999
6.1335 | -102.154
6.1373 | -102.109
6.1411 | -102.0725
6.1448 | -102.0475
6.1486 | -102.0332
6.1524 | -102.0285
6.1562 | -102.0302
6.1599 | -102.047
6.1637 | -102.0774
6.1675 | -102.1047
6.1713 | -102.132
6.175 | -102.1543
6.1788 | -102.1708
6.1826 | -102.1697
6.1864 | -102.1568
6.1901 | -102.1364
6.1939 | -102.1052
6.1977 | -102.0679
6.2015 | -102.0392
6.2052 | -102.0249
6.209 | -102.0186
6.2128 | -102.025
6.2166 | -102.0433
6.2203 | -102.0747
6.2241 | -102.1214
6.2279 | -102.1735
6.2317 | -102.2261
6.2354 | -102.2818
6.2392 | -102.3394
6.243 | -102.3957
6.2468 | -102.4445
6.2505 | -102.4806
6.2543 | -102.4946
6.2581 | -102.4984
6.2619 | -102.4945
6.2656 | -102.4735
6.2694 | -102.4345
6.2732 | -102.3892
6.277 | -102.3468
6.2807 | -102.3023
6.2845 | -102.2685
6.2883 | -102.2478
6.292 | -102.232
6.2958 | -102.2245
6.2996 | -102.2312
6.3034 | -102.2483
6.3071 | -102.269
6.3109 | -102.2887
6.3147 | -102.3049
6.3185 | -102.3179
6.3222 | -102.3279
6.326 | -102.3367
6.3298 | -102.3381
6.3336 | -102.3347
6.3373 | -102.3272
6.3411 | -102.3181
6.3449 | -102.3094
6.3487 | -102.2956
6.3524 | -102.2693
6.3562 | -102.245
6.36 | -102.231
6.3638 | -102.2241
6.3675 | -102.2164
6.3713 | -102.2092
6.3751 | -102.213
6.3789 | -102.2222
6.3826 | -102.2323
6.3864 | -102.2465
6.3902 | -102.2513
6.394 | -102.2426
6.3977 | -102.2327
6.4015 | -102.2215
6.4053 | -102.2028
6.4091 | -102.1827
6.4128 | -102.1618
6.4166 | -102.1347
6.4204 | -102.1067
6.4242 | -102.0829
6.4279 | -102.064
6.4317 | -102.0498
6.4355 | -102.0404
6.4393 | -102.0419
6.443 | -102.0503
6.4468 | -102.0606
6.4506 | -102.0681
6.4544 | -102.068
6.4581 | -102.0742
6.4619 | -102.0885
6.4657 | -102.1063
6.4695 | -102.1232
6.4732 | -102.1538
6.477 | -102.1932
6.4808 | -102.242
6.4845 | -102.3111
6.4883 | -102.3922
6.4921 | -102.4698
6.4959 | -102.5491
6.4996 | -102.6253
6.5034 | -102.6898
6.5072 | -102.7389
6.511 | -102.78
6.5147 | -102.8023
6.5185 | -102.8036
6.5223 | -102.7958
6.5261 | -102.7778
6.5298 | -102.7524
6.5336 | -102.7108
6.5374 | -102.6768
6.5412 | -102.6427
6.5449 | -102.6066
6.5487 | -102.5602
6.5525 | -102.514
6.5563 | -102.4666
6.56 | -102.4204
6.5638 | -102.3801
6.5676 | -102.3493
6.5714 | -102.3197
6.5751 | -102.3007
6.5789 | -102.2875
6.5827 | -102.2801
6.5865 | -102.2819
6.5902 | -102.296
6.594 | -102.3166
6.5978 | -102.3435
6.6016 | -102.3719
6.6053 | -102.4036
6.6091 | -102.4346
6.6129 | -102.4604
6.6167 | -102.4822
6.6204 | -102.4953
6.6242 | -102.5088
6.628 | -102.5154
6.6318 | -102.5189
6.6355 | -102.5183
6.6393 | -102.5141
6.6431 | -102.51
6.6469 | -102.5142
6.6506 | -102.5182
6.6544 | -102.5167
6.6582 | -102.5146
6.6619 | -102.5105
6.6657 | -102.4968
6.6695 | -102.4806
6.6733 | -102.4671
6.677 | -102.4509
6.6808 | -102.4374
6.6846 | -102.4234
6.6884 | -102.4087
6.6921 | -102.3961
6.6959 | -102.3902
6.6997 | -102.3954
6.7035 | -102.4067
6.7072 | -102.4232
6.711 | -102.4485
6.7148 | -102.4837
6.7186 | -102.5236
6.7223 | -102.5686
6.7261 | -102.6146
6.7299 | -102.6677
6.7337 | -102.7188
6.7374 | -102.7781
6.7412 | -102.8428
6.745 | -102.9035
6.7488 | -102.9553
6.7525 | -103.0086
6.7563 | -103.0599
6.7601 | -103.0932
6.7639 | -103.1166
6.7676 | -103.1283
6.7714 | -103.1144
6.7752 | -103.0813
6.779 | -103.0356
6.7827 | -102.977
6.7865 | -102.9098
6.7903 | -102.844
6.7941 | -102.782
6.7978 | -102.7254
6.8016 | -102.6771
6.8054 | -102.6372
6.8092 | -102.6087
6.8129 | -102.5962
6.8167 | -102.5952
6.8205 | -102.5973
6.8243 | -102.6028
6.828 | -102.6071
6.8318 | -102.6056
6.8356 | -102.6066
6.8393 | -102.6057
6.8431 | -102.5933
6.8469 | -102.5816
6.8507 | -102.5645
6.8544 | -102.5422
6.8582 | -102.5138
6.862 | -102.4851
6.8658 | -102.4578
6.8695 | -102.434
6.8733 | -102.417
6.8771 | -102.4014
6.8809 | -102.3986
6.8846 | -102.4079
6.8884 | -102.431
6.8922 | -102.466
6.896 | -102.5112
6.8997 | -102.5603
6.9035 | -102.6096
6.9073 | -102.6659
6.9111 | -102.7171
6.9148 | -102.7528
6.9186 | -102.7777
6.9224 | -102.7937
6.9262 | -102.7953
6.9299 | -102.7879
6.9337 | -102.773
6.9375 | -102.7418
6.9413 | -102.6953
6.945 | -102.6514
6.9488 | -102.6065
6.9526 | -102.5621
6.9564 | -102.5205
6.9601 | -102.4865
6.9639 | -102.4664
6.9677 | -102.4538
6.9715 | -102.4574
6.9752 | -102.4698
6.979 | -102.4901
6.9828 | -102.5177
6.9866 | -102.5467
6.9903 | -102.5781
6.9941 | -102.6157
6.9979 | -102.658
7.0017 | -102.7041
7.0054 | -102.7561
7.0092 | -102.8056
7.013 | -102.8558
7.0167 | -102.9082
7.0205 | -102.9606
7.0243 | -103.0055
7.0281 | -103.044
7.0318 | -103.0797
7.0356 | -103.1085
7.0394 | -103.1374
7.0432 | -103.1664
7.0469 | -103.1972
7.0507 | -103.2287
7.0545 | -103.2633
7.0583 | -103.2957
7.062 | -103.3192
7.0658 | -103.3322
7.0696 | -103.3444
7.0734 | -103.3505
7.0771 | -103.3416
7.0809 | -103.3357
7.0847 | -103.3352
7.0885 | -103.3266
7.0922 | -103.3179
7.096 | -103.3086
7.0998 | -103.2981
7.1036 | -103.2875
7.1073 | -103.2808
7.1111 | -103.2725
7.1149 | -103.262
7.1187 | -103.2535
7.1224 | -103.2482
7.1262 | -103.2427
7.13 | -103.2387
7.1338 | -103.2323
7.1375 | -103.2218
7.1413 | -103.2139
7.1451 | -103.2034
7.1489 | -103.1887
7.1526 | -103.1738
7.1564 | -103.1603
7.1602 | -103.1481
7.164 | -103.1317
7.1677 | -103.1233
7.1715 | -103.115
7.1753 | -103.1024
7.1791 | -103.094
7.1828 | -103.0916
7.1866 | -103.102
7.1904 | -103.1167
7.1941 | -103.1233
7.1979 | -103.1414
7.2017 | -103.1707
7.2055 | -103.1995
7.2092 | -103.2346
7.213 | -103.2725
7.2168 | -103.3092
7.2206 | -103.3444
7.2243 | -103.3815
7.2281 | -103.4156
7.2319 | -103.4441
7.2357 | -103.4633
7.2394 | -103.4743
7.2432 | -103.4827
7.247 | -103.4875
7.2508 | -103.49
7.2545 | -103.4912
7.2583 | -103.4893
7.2621 | -103.4756
7.2659 | -103.4616
7.2696 | -103.4464
7.2734 | -103.4353
7.2772 | -103.4183
7.281 | -103.4076
7.2847 | -103.3952
7.2885 | -103.3815
7.2923 | -103.3711
7.2961 | -103.3558
7.2998 | -103.3455
7.3036 | -103.3421
7.3074 | -103.3344
7.3112 | -103.3241
7.3149 | -103.3175
7.3187 | -103.3047
7.3225 | -103.2969
7.3263 | -103.2938
7.33 | -103.2989
7.3338 | -103.3064
7.3376 | -103.3184
7.3414 | -103.3424
7.3451 | -103.3774
7.3489 | -103.4152
7.3527 | -103.4419
7.3565 | -103.4623
7.3602 | -103.4775
7.364 | -103.4864
7.3678 | -103.4903
7.3716 | -103.4884
7.3753 | -103.4702
7.3791 | -103.4433
7.3829 | -103.4085
7.3866 | -103.376
7.3904 | -103.3399
7.3942 | -103.3078
7.398 | -103.2776
7.4017 | -103.2508
7.4055 | -103.2357
7.4093 | -103.2253
7.4131 | -103.2254
7.4168 | -103.2303
7.4206 | -103.243
7.4244 | -103.2621
7.4282 | -103.2851
7.4319 | -103.3047
7.4357 | -103.3246
7.4395 | -103.3432
7.4433 | -103.3558
7.447 | -103.3543
7.4508 | -103.3504
7.4546 | -103.3418
7.4584 | -103.332
7.4621 | -103.3212
7.4659 | -103.3056
7.4697 | -103.2825
7.4735 | -103.2581
7.4772 | -103.235
7.481 | -103.2181
7.4848 | -103.2051
7.4886 | -103.19
7.4923 | -103.1801
7.4961 | -103.1682
7.4999 | -103.1617
7.5037 | -103.1624
7.5074 | -103.1706
7.5112 | -103.1849
7.515 | -103.2009
7.5188 | -103.2189
7.5225 | -103.2373
7.5263 | -103.2573
7.5301 | -103.2768
7.5339 | -103.2953
7.5376 | -103.3139
7.5414 | -103.3364
7.5452 | -103.3577
7.5489 | -103.3787
7.5527 | -103.402
7.5565 | -103.4319
7.5603 | -103.4543
7.564 | -103.4787
7.5678 | -103.5004
7.5716 | -103.5262
7.5754 | -103.5479
7.5791 | -103.5702
7.5829 | -103.5862
7.5867 | -103.5991
7.5905 | -103.6098
7.5942 | -103.617
7.598 | -103.627
7.6018 | -103.633
7.6056 | -103.6393
7.6093 | -103.6353
7.6131 | -103.6246
7.6169 | -103.6079
7.6207 | -103.5538
7.6244 | -103.5168
7.6282 | -103.4705
7.632 | -103.4123
7.6358 | -103.3534
7.6395 | -103.295
7.6433 | -103.2347
7.6471 | -103.1738
7.6509 | -103.1163
7.6546 | -103.0651
7.6584 | -103.0219
7.6622 | -102.9885
7.666 | -102.9655
7.6697 | -102.953
7.6735 | -102.9533
7.6773 | -102.9537
7.6811 | -102.9624
7.6848 | -102.9722
7.6886 | -102.9827
7.6924 | -102.9946
7.6962 | -103.0119
7.6999 | -103.0229
7.7037 | -103.0297
7.7075 | -103.0353
7.7113 | -103.0358
7.715 | -103.0359
7.7188 | -103.0278
7.7226 | -103.0235
7.7264 | -103.0224
7.7301 | -103.0161
7.7339 | -103.0152
7.7377 | -103.0129
7.7414 | -103.0174
7.7452 | -103.0322
7.749 | -103.0542
7.7528 | -103.0777
7.7565 | -103.107
7.7603 | -103.146
7.7641 | -103.1984
7.7679 | -103.2615
7.7716 | -103.3282
7.7754 | -103.4032
7.7792 | -103.4816
7.783 | -103.5615
7.7867 | -103.6385
7.7905 | -103.7043
7.7943 | -103.7704
7.7981 | -103.8245
7.8018 | -103.866
7.8056 | -103.8946
7.8094 | -103.9076
7.8132 | -103.9032
7.8169 | -103.893
7.8207 | -103.8745
7.8245 | -103.8492
7.8283 | -103.8235
7.832 | -103.7935
7.8358 | -103.767
7.8396 | -103.7367
7.8434 | -103.704
7.8471 | -103.6765
7.8509 | -103.6496
7.8547 | -103.6264
7.8585 | -103.6046
7.8622 | -103.5855
7.866 | -103.5672
7.8698 | -103.5511
7.8736 | -103.5359
7.8773 | -103.5146
7.8811 | -103.4901
7.8849 | -103.4658
7.8887 | -103.4416
7.8924 | -103.4185
7.8962 | -103.4002
7.9 | -103.3844
7.9038 | -103.372
7.9075 | -103.3697
7.9113 | -103.3741
7.9151 | -103.3847
7.9188 | -103.4009
7.9226 | -103.4203
7.9264 | -103.4412
7.9302 | -103.4657
7.9339 | -103.4874
7.9377 | -103.5027
7.9415 | -103.514
7.9453 | -103.5183
7.949 | -103.5189
7.9528 | -103.5198
7.9566 | -103.5239
7.9604 | -103.5295
7.9641 | -103.5369
7.9679 | -103.547
7.9717 | -103.555
7.9755 | -103.5605
7.9792 | -103.5668
7.983 | -103.5823
7.9868 | -103.5957
7.9906 | -103.6061
7.9943 | -103.6236
7.9981 | -103.6398
8.0019 | -103.6549
8.0057 | -103.676
8.0094 | -103.6962
8.0132 | -103.7107
8.017 | -103.725
8.0208 | -103.7376
8.0245 | -103.7423
8.0283 | -103.7396
8.0321 | -103.7347
8.0359 | -103.7225
8.0396 | -103.7074
8.0434 | -103.6942
8.0472 | -103.6793
8.051 | -103.6647
8.0547 | -103.6536
8.0585 | -103.6453
8.0623 | -103.6382
8.0661 | -103.6352
8.0698 | -103.6361
8.0736 | -103.635
8.0774 | -103.6377
8.0812 | -103.6442
8.0849 | -103.6561
8.0887 | -103.6728
8.0925 | -103.6868
8.0962 | -103.7041
8.1 | -103.7183
8.1038 | -103.7337
8.1076 | -103.7492
8.1113 | -103.7578
8.1151 | -103.7588
8.1189 | -103.7568
8.1227 | -103.7526
8.1264 | -103.7447
8.1302 | -103.7369
8.134 | -103.728
8.1378 | -103.7182
8.1415 | -103.7123
8.1453 | -103.7108
8.1491 | -103.7111
8.1529 | -103.7202
8.1566 | -103.7314
8.1604 | -103.7384
8.1642 | -103.7519
8.168 | -103.7859
8.1717 | -103.81
8.1755 | -103.8353
8.1793 | -103.8688
8.1831 | -103.902
8.1868 | -103.9339
8.1906 | -103.9641
8.1944 | -104.0005
8.1982 | -104.0265
8.2019 | -104.0451
8.2057 | -104.0611
8.2095 | -104.0708
8.2133 | -104.0735
8.217 | -104.0664
8.2208 | -104.0585
8.2246 | -104.0464
8.2284 | -104.0239
8.2321 | -103.9961
8.2359 | -103.9656
8.2397 | -103.9254
8.2435 | -103.884
8.2472 | -103.8439
8.251 | -103.8051
8.2548 | -103.7613
8.2586 | -103.7141
8.2623 | -103.675
8.2661 | -103.6391
8.2699 | -103.6079
8.2736 | -103.5857
8.2774 | -103.5715
8.2812 | -103.5666
8.285 | -103.5663
8.2887 | -103.5657
8.2925 | -103.5699
8.2963 | -103.5817
8.3001 | -103.6018
8.3038 | -103.6203
8.3076 | -103.6433
8.3114 | -103.6677
8.3152 | -103.6982
8.3189 | -103.7288
8.3227 | -103.7682
8.3265 | -103.8118
8.3303 | -103.8556
8.334 | -103.9069
8.3378 | -103.9623
8.3416 | -104.0188
8.3454 | -104.072
8.3491 | -104.1256
8.3529 | -104.1885
8.3567 | -104.2443
8.3605 | -104.293
8.3642 | -104.3384
8.368 | -104.3736
8.3718 | -104.4041
8.3756 | -104.4274
8.3793 | -104.4448
8.3831 | -104.4542
8.3869 | -104.4469
8.3907 | -104.4368
8.3944 | -104.4199
8.3982 | -104.3967
8.402 | -104.3711
8.4058 | -104.3442
8.4095 | -104.3147
8.4133 | -104.284
8.4171 | -104.2498
8.4209 | -104.2149
8.4246 | -104.1802
8.4284 | -104.1486
8.4322 | -104.1241
8.436 | -104.1034
8.4397 | -104.0837
8.4435 | -104.0643
8.4473 | -104.0515
8.4511 | -104.041
8.4548 | -104.0246
8.4586 | -104.0028
8.4624 | -103.9832
8.4661 | -103.9691
8.4699 | -103.9593
8.4737 | -103.9539
8.4775 | -103.9433
8.4812 | -103.9361
8.485 | -103.9337
8.4888 | -103.9254
8.4926 | -103.9172
8.4963 | -103.9062
8.5001 | -103.8947
8.5039 | -103.8849
8.5077 | -103.87
8.5114 | -103.8558
8.5152 | -103.8357
8.519 | -103.8153
8.5228 | -103.7965
8.5265 | -103.7777
8.5303 | -103.76
8.5341 | -103.7504
8.5379 | -103.7443
8.5416 | -103.74
8.5454 | -103.733
8.5492 | -103.7263
8.553 | -103.721
8.5567 | -103.7198
8.5605 | -103.723
8.5643 | -103.7305
8.5681 | -103.7459
8.5718 | -103.7636
8.5756 | -103.7854
8.5794 | -103.8184
8.5832 | -103.8606
8.5869 | -103.9064
8.5907 | -103.9508
8.5945 | -103.9975
8.5983 | -104.046
8.602 | -104.1007
8.6058 | -104.16
8.6096 | -104.2162
8.6134 | -104.2727
8.6171 | -104.3233
8.6209 | -104.3731
8.6247 | -104.4223
8.6285 | -104.4665
8.6322 | -104.5136
8.636 | -104.5542
8.6398 | -104.5898
8.6435 | -104.6246
8.6473 | -104.6493
8.6511 | -104.6654
8.6549 | -104.6805
8.6586 | -104.6891
8.6624 | -104.695
8.6662 | -104.6928
8.67 | -104.686
8.6737 | -104.6729
8.6775 | -104.6535
8.6813 | -104.6269
8.6851 | -104.5969
8.6888 | -104.5581
8.6926 | -104.5115
8.6964 | -104.4645
8.7002 | -104.4223
8.7039 | -104.3799
8.7077 | -104.3344
8.7115 | -104.3048
8.7153 | -104.273
8.719 | -104.2435
8.7228 | -104.2206
8.7266 | -104.2029
8.7304 | -104.1912
8.7341 | -104.1873
8.7379 | -104.19
8.7417 | -104.1961
8.7455 | -104.2012
8.7492 | -104.2112
8.753 | -104.2225
8.7568 | -104.2359
8.7606 | -104.2466
8.7643 | -104.2566
8.7681 | -104.2729
8.7719 | -104.2881
8.7757 | -104.3007
8.7794 | -104.3134
8.7832 | -104.3216
8.787 | -104.3259
8.7908 | -104.332
8.7945 | -104.3389
8.7983 | -104.3456
8.8021 | -104.3506
8.8058 | -104.3508
8.8096 | -104.3468
8.8134 | -104.3444
8.8172 | -104.3485
8.8209 | -104.3553
8.8247 | -104.356
8.8285 | -104.3525
8.8323 | -104.3435
8.836 | -104.3362
8.8398 | -104.3271
8.8436 | -104.3159
8.8474 | -104.302
8.8511 | -104.2869
8.8549 | -104.2697
8.8587 | -104.2568
8.8625 | -104.244
8.8662 | -104.2257
8.87 | -104.2147
8.8738 | -104.2154
8.8776 | -104.2227
8.8813 | -104.2263
8.8851 | -104.2317
8.8889 | -104.2406
8.8927 | -104.2521
8.8964 | -104.2624
8.9002 | -104.269
8.904 | -104.2756
8.9078 | -104.2843
8.9115 | -104.2965
8.9153 | -104.3072
8.9191 | -104.3189
8.9229 | -104.3308
8.9266 | -104.3412
8.9304 | -104.3506
8.9342 | -104.3611
8.938 | -104.3744
8.9417 | -104.3858
8.9455 | -104.3918
8.9493 | -104.3972
8.9531 | -104.4023
8.9568 | -104.4091
8.9606 | -104.4147
8.9644 | -104.4203
8.9682 | -104.4262
8.9719 | -104.4306
8.9757 | -104.4328
8.9795 | -104.4345
8.9833 | -104.4393
8.987 | -104.445
8.9908 | -104.446
8.9946 | -104.4491
8.9983 | -104.4574
9.0021 | -104.463
9.0059 | -104.4667
9.0097 | -104.4717
9.0134 | -104.4774
9.0172 | -104.4857
9.021 | -104.4934
9.0248 | -104.5024
9.0285 | -104.5107
9.0323 | -104.515
9.0361 | -104.5155
9.0399 | -104.5182
9.0436 | -104.5124
9.0474 | -104.5069
9.0512 | -104.5057
9.055 | -104.5045
9.0587 | -104.506
9.0625 | -104.5072
9.0663 | -104.5111
9.0701 | -104.5153
9.0738 | -104.5236
9.0776 | -104.5343
9.0814 | -104.5421
9.0852 | -104.5502
9.0889 | -104.5612
9.0927 | -104.5724
9.0965 | -104.5867
9.1003 | -104.598
9.104 | -104.6067
9.1078 | -104.6164
9.1116 | -104.6302
9.1154 | -104.6382
9.1191 | -104.6358
9.1229 | -104.6348
9.1267 | -104.6319
9.1305 | -104.6228
9.1342 | -104.6098
9.138 | -104.5903
9.1418 | -104.5712
9.1456 | -104.5454
9.1493 | -104.5165
9.1531 | -104.4891
9.1569 | -104.4601
9.1607 | -104.4331
9.1644 | -104.409
9.1682 | -104.3848
9.172 | -104.3595
9.1757 | -104.3421
9.1795 | -104.323
9.1833 | -104.2998
9.1871 | -104.2797
9.1908 | -104.2678
9.1946 | -104.2597
9.1984 | -104.2549
9.2022 | -104.2538
9.2059 | -104.257
9.2097 | -104.2619
9.2135 | -104.2703
9.2173 | -104.2859
9.221 | -104.301
9.2248 | -104.3189
9.2286 | -104.3411
9.2324 | -104.3625
9.2361 | -104.3852
9.2399 | -104.4065
9.2437 | -104.4193
9.2475 | -104.4324
9.2512 | -104.4494
9.255 | -104.4497
9.2588 | -104.4568
9.2626 | -104.4588
9.2663 | -104.4633
9.2701 | -104.4659
9.2739 | -104.4699
9.2777 | -104.4735
9.2814 | -104.4827
9.2852 | -104.493
9.289 | -104.4999
9.2928 | -104.5098
9.2965 | -104.5199
9.3003 | -104.5319
9.3041 | -104.547
9.3079 | -104.5596
9.3116 | -104.5712
9.3154 | -104.5826
9.3192 | -104.5972
9.323 | -104.605
9.3267 | -104.6047
9.3305 | -104.6072
9.3343 | -104.6091
9.3381 | -104.607
9.3418 | -104.6039
9.3456 | -104.5984
9.3494 | -104.5915
9.3531 | -104.5782
9.3569 | -104.5628
9.3607 | -104.5433
9.3645 | -104.5177
9.3682 | -104.4989
9.372 | -104.4836
9.3758 | -104.4692
9.3796 | -104.4624
9.3833 | -104.4608
9.3871 | -104.4653
9.3909 | -104.477
9.3947 | -104.4876
9.3984 | -104.504
9.4022 | -104.5281
9.406 | -104.5527
9.4098 | -104.578
9.4135 | -104.6058
9.4173 | -104.6309
9.4211 | -104.6518
9.4249 | -104.6661
9.4286 | -104.675
9.4324 | -104.6775
9.4362 | -104.676
9.44 | -104.6619
9.4437 | -104.6389
9.4475 | -104.6095
9.4513 | -104.5754
9.4551 | -104.5417
9.4588 | -104.5024
9.4626 | -104.4652
9.4664 | -104.4276
9.4702 | -104.3873
9.4739 | -104.3473
9.4777 | -104.3137
9.4815 | -104.2833
9.4853 | -104.2605
9.489 | -104.2456
9.4928 | -104.2359
9.4966 | -104.2307
9.5004 | -104.2327
9.5041 | -104.243
9.5079 | -104.2509
9.5117 | -104.2643
9.5155 | -104.2847
9.5192 | -104.308
9.523 | -104.3367
9.5268 | -104.3697
9.5305 | -104.4065
9.5343 | -104.45
9.5381 | -104.4998
9.5419 | -104.5494
9.5456 | -104.6041
9.5494 | -104.6642
9.5532 | -104.7217
9.557 | -104.7798
9.5607 | -104.8335
9.5645 | -104.8847
9.5683 | -104.9404
9.5721 | -104.9913
9.5758 | -105.0328
9.5796 | -105.0729
9.5834 | -105.1039
9.5872 | -105.1279
9.5909 | -105.145
9.5947 | -105.1505
9.5985 | -105.1494
9.6023 | -105.1461
9.606 | -105.1377
9.6098 | -105.1222
9.6136 | -105.1054
9.6174 | -105.0875
9.6211 | -105.0612
9.6249 | -105.0313
9.6287 | -104.9984
9.6325 | -104.9641
9.6362 | -104.9296
9.64 | -104.8947
9.6438 | -104.8611
9.6476 | -104.8294
9.6513 | -104.7944
9.6551 | -104.7563
9.6589 | -104.7221
9.6627 | -104.6862
9.6664 | -104.6531
9.6702 | -104.6232
9.674 | -104.5926
9.6778 | -104.5618
9.6815 | -104.5369
9.6853 | -104.5145
9.6891 | -104.4932
9.6929 | -104.4767
9.6966 | -104.463
9.7004 | -104.4541
9.7042 | -104.4455
9.708 | -104.4323
9.7117 | -104.4237
9.7155 | -104.4224
9.7193 | -104.4241
9.723 | -104.4291
9.7268 | -104.4319
9.7306 | -104.4334
9.7344 | -104.4357
9.7381 | -104.441
9.7419 | -104.4452
9.7457 | -104.4527
9.7495 | -104.4639
9.7532 | -104.4725
9.757 | -104.4806
9.7608 | -104.49
9.7646 | -104.4984
9.7683 | -104.5077
9.7721 | -104.5173
9.7759 | -104.529
9.7797 | -104.5406
9.7834 | -104.5518
9.7872 | -104.5685
9.791 | -104.5884
9.7948 | -104.6118
9.7985 | -104.64
9.8023 | -104.6649
9.8061 | -104.69
9.8099 | -104.7132
9.8136 | -104.7398
9.8174 | -104.7687
9.8212 | -104.7954
9.825 | -104.8189
9.8287 | -104.8404
9.8325 | -104.8638
9.8363 | -104.8871
9.8401 | -104.9064
9.8438 | -104.9221
9.8476 | -104.9354
9.8514 | -104.9463
9.8552 | -104.9512
9.8589 | -104.947
9.8627 | -104.9417
9.8665 | -104.9355
9.8703 | -104.9317
9.874 | -104.9222
9.8778 | -104.9107
9.8816 | -104.8971
9.8854 | -104.8799
9.8891 | -104.8614
9.8929 | -104.8427
9.8967 | -104.8269
9.9004 | -104.8106
9.9042 | -104.7976
9.908 | -104.7857
9.9118 | -104.7769
9.9155 | -104.7704
9.9193 | -104.7627
9.9231 | -104.7559
9.9269 | -104.748
9.9306 | -104.7401
9.9344 | -104.7351
9.9382 | -104.7305
9.942 | -104.7293
9.9457 | -104.7299
9.9495 | -104.7276
9.9533 | -104.73
9.9571 | -104.7365
9.9608 | -104.7431
9.9646 | -104.753
9.9684 | -104.7653
9.9722 | -104.78
9.9759 | -104.7993
9.9797 | -104.8193
9.9835 | -104.8407
9.9873 | -104.8682
9.991 | -104.8943
9.9948 | -104.9199
9.9986 | -104.9493
10.0024 | -104.9761
10.0061 | -105.0034
10.0099 | -105.0243
10.0137 | -105.0465
10.0175 | -105.0667
10.0212 | -105.0896
10.025 | -105.1114
10.0288 | -105.1317
10.0326 | -105.154
10.0363 | -105.173
10.0401 | -105.188
10.0439 | -105.201
10.0477 | -105.2125
10.0514 | -105.2205
10.0552 | -105.2218
10.059 | -105.2113
10.0628 | -105.2004
10.0665 | -105.1886
10.0703 | -105.1716
10.0741 | -105.1555
10.0778 | -105.1361
10.0816 | -105.1095
10.0854 | -105.0802
10.0892 | -105.0474
10.0929 | -105.0139
10.0967 | -104.983
10.1005 | -104.9504
10.1043 | -104.9168
10.108 | -104.8856
10.1118 | -104.8554
10.1156 | -104.8304
10.1194 | -104.8115
10.1231 | -104.795
10.1269 | -104.7842
10.1307 | -104.7791
10.1345 | -104.7768
10.1382 | -104.7839
10.142 | -104.795
10.1458 | -104.8074
10.1496 | -104.8276
10.1533 | -104.85
10.1571 | -104.8732
10.1609 | -104.8887
10.1647 | -104.9019
10.1684 | -104.9155
10.1722 | -104.9276
10.176 | -104.9358
10.1798 | -104.938
10.1835 | -104.9379
10.1873 | -104.9371
10.1911 | -104.9348
10.1949 | -104.9286
10.1986 | -104.9163
10.2024 | -104.9045
10.2062 | -104.8979
10.21 | -104.8869
10.2137 | -104.8754
10.2175 | -104.8628
10.2213 | -104.8523
10.2251 | -104.8439
10.2288 | -104.8403
10.2326 | -104.8405
10.2364 | -104.8451
10.2402 | -104.8499
10.2439 | -104.8539
10.2477 | -104.8591
10.2515 | -104.8674
10.2552 | -104.8777
10.259 | -104.891
10.2628 | -104.9048
10.2666 | -104.9143
10.2703 | -104.9219
10.2741 | -104.9311
10.2779 | -104.9417
10.2817 | -104.9496
10.2854 | -104.9509
10.2892 | -104.9468
10.293 | -104.9415
10.2968 | -104.933
10.3005 | -104.9252
10.3043 | -104.9166
10.3081 | -104.9086
10.3119 | -104.8977
10.3156 | -104.8838
10.3194 | -104.8685
10.3232 | -104.8496
10.327 | -104.8302
10.3307 | -104.8133
10.3345 | -104.7964
10.3383 | -104.7816
10.3421 | -104.7702
10.3458 | -104.7704
10.3496 | -104.7645
10.3534 | -104.7586
10.3572 | -104.7523
10.3609 | -104.7506
10.3647 | -104.7476
10.3685 | -104.749
10.3723 | -104.7477
10.376 | -104.7456
10.3798 | -104.7467
10.3836 | -104.7485
10.3874 | -104.7521
10.3911 | -104.7565
10.3949 | -104.7635
10.3987 | -104.7694
10.4025 | -104.7739
10.4062 | -104.7799
10.41 | -104.7833
10.4138 | -104.7884
10.4176 | -104.8004
10.4213 | -104.8129
10.4251 | -104.8156
10.4289 | -104.8145
10.4326 | -104.8175
10.4364 | -104.827
10.4402 | -104.8355
10.444 | -104.8373
10.4477 | -104.8395
10.4515 | -104.8362
10.4553 | -104.8314
10.4591 | -104.8256
10.4628 | -104.8215
10.4666 | -104.8151
10.4704 | -104.8084
10.4742 | -104.8068
10.4779 | -104.8086
10.4817 | -104.809
10.4855 | -104.8079
10.4893 | -104.8091
10.493 | -104.8148
10.4968 | -104.8229
10.5006 | -104.8344
10.5044 | -104.8505
10.5081 | -104.8685
10.5119 | -104.8916
10.5157 | -104.9169
10.5195 | -104.9482
10.5232 | -104.9803
10.527 | -105.0175
10.5308 | -105.0536
10.5346 | -105.0868
10.5383 | -105.123
10.5421 | -105.157
10.5459 | -105.192
10.5497 | -105.2244
10.5534 | -105.2507
10.5572 | -105.2748
10.561 | -105.2936
10.5648 | -105.3085
10.5685 | -105.3179
10.5723 | -105.318
10.5761 | -105.3148
10.5799 | -105.3111
10.5836 | -105.3007
10.5874 | -105.2824
10.5912 | -105.2578
10.595 | -105.2346
10.5987 | -105.2143
10.6025 | -105.1889
10.6063 | -105.1679
10.61 | -105.1458
10.6138 | -105.1246
10.6176 | -105.1037
10.6214 | -105.0844
10.6251 | -105.0684
10.6289 | -105.0513
10.6327 | -105.0389
10.6365 | -105.0302
10.6402 | -105.0247
10.644 | -105.021
10.6478 | -105.0181
10.6516 | -105.0202
10.6553 | -105.0217
10.6591 | -105.0262
10.6629 | -105.0324
10.6667 | -105.0372
10.6704 | -105.0449
10.6742 | -105.0529
10.678 | -105.0614
10.6818 | -105.0686
10.6855 | -105.0812
10.6893 | -105.0961
10.6931 | -105.1086
10.6969 | -105.1245
10.7006 | -105.1431
10.7044 | -105.1621
10.7082 | -105.1842
10.712 | -105.2053
10.7157 | -105.2255
10.7195 | -105.2461
10.7233 | -105.2673
10.7271 | -105.29
10.7308 | -105.3127
10.7346 | -105.335
10.7384 | -105.3565
10.7422 | -105.3765
10.7459 | -105.3941
10.7497 | -105.4091
10.7535 | -105.4174
10.7573 | -105.4279
10.761 | -105.4288
10.7648 | -105.4315
10.7686 | -105.434
10.7724 | -105.4294
10.7761 | -105.4232
10.7799 | -105.4175
10.7837 | -105.4117
10.7874 | -105.398
10.7912 | -105.3842
10.795 | -105.3723
10.7988 | -105.3554
10.8025 | -105.3396
10.8063 | -105.3253
10.8101 | -105.3147
10.8139 | -105.3053
10.8176 | -105.2969
10.8214 | -105.2886
10.8252 | -105.2842
10.829 | -105.2826
10.8327 | -105.286
10.8365 | -105.2917
10.8403 | -105.2976
10.8441 | -105.3017
10.8478 | -105.3065
10.8516 | -105.3133
10.8554 | -105.3209
10.8592 | -105.3277
10.8629 | -105.3354
10.8667 | -105.3405
10.8705 | -105.3457
10.8743 | -105.3474
10.878 | -105.3408
10.8818 | -105.3309
10.8856 | -105.3191
10.8894 | -105.2902
10.8931 | -105.273
10.8969 | -105.2531
10.9007 | -105.2296
10.9045 | -105.2033
10.9082 | -105.1738
10.912 | -105.1446
10.9158 | -105.1135
10.9196 | -105.0846
10.9233 | -105.0529
10.9271 | -105.0219
10.9309 | -104.9941
10.9347 | -104.9671
10.9384 | -104.9408
10.9422 | -104.9188
10.946 | -104.9043
10.9498 | -104.8906
10.9535 | -104.8777
10.9573 | -104.8756
10.9611 | -104.875
10.9649 | -104.8753
10.9686 | -104.8801
10.9724 | -104.8874
10.9762 | -104.8957
10.9799 | -104.9065
10.9837 | -104.9163
10.9875 | -104.9305
10.9913 | -104.9469
10.995 | -104.9637
10.9988 | -104.9823
11.0026 | -105.0022
11.0064 | -105.0263
11.0101 | -105.049
11.0139 | -105.0721
11.0177 | -105.098
11.0215 | -105.1242
11.0252 | -105.1486
11.029 | -105.1706
11.0328 | -105.1949
11.0366 | -105.2216
11.0403 | -105.2454
11.0441 | -105.264
11.0479 | -105.2783
11.0517 | -105.2912
11.0554 | -105.3007
11.0592 | -105.307
11.063 | -105.3096
11.0668 | -105.3084
11.0705 | -105.3072
11.0743 | -105.3045
11.0781 | -105.2968
11.0819 | -105.287
11.0856 | -105.2738
11.0894 | -105.2559
11.0932 | -105.2365
11.097 | -105.2165
11.1007 | -105.1971
11.1045 | -105.175
11.1083 | -105.1524
11.1121 | -105.1354
11.1158 | -105.1204
11.1196 | -105.1038
11.1234 | -105.0931
11.1272 | -105.0846
11.1309 | -105.0821
11.1347 | -105.0835
11.1385 | -105.0892
11.1423 | -105.0981
11.146 | -105.1117
11.1498 | -105.1274
11.1536 | -105.1446
11.1573 | -105.1668
11.1611 | -105.1911
11.1649 | -105.2185
11.1687 | -105.2453
11.1724 | -105.2738
11.1762 | -105.3003
11.18 | -105.3246
11.1838 | -105.3501
11.1875 | -105.3762
11.1913 | -105.3964
11.1951 | -105.4128
11.1989 | -105.428
11.2026 | -105.4402
11.2064 | -105.4503
11.2102 | -105.4585
11.214 | -105.4611
11.2177 | -105.4622
11.2215 | -105.4571
11.2253 | -105.4516
11.2291 | -105.446
11.2328 | -105.4354
11.2366 | -105.419
11.2404 | -105.4047
11.2442 | -105.39
11.2479 | -105.3726
11.2517 | -105.3544
11.2555 | -105.3409
11.2593 | -105.3275
11.263 | -105.3101
11.2668 | -105.2979
11.2706 | -105.2884
11.2744 | -105.2738
11.2781 | -105.2623
11.2819 | -105.2519
11.2857 | -105.2432
11.2895 | -105.2376
11.2932 | -105.2268
11.297 | -105.2181
11.3008 | -105.2098
11.3046 | -105.1994
11.3083 | -105.1915
11.3121 | -105.1785
11.3159 | -105.1662
11.3197 | -105.1559
11.3234 | -105.1469
11.3272 | -105.1389
11.331 | -105.131
11.3347 | -105.1239
11.3385 | -105.1194
11.3423 | -105.1164
11.3461 | -105.1139
11.3498 | -105.1086
11.3536 | -105.1073
11.3574 | -105.109
11.3612 | -105.1071
11.3649 | -105.1078
11.3687 | -105.1106
11.3725 | -105.1118
11.3763 | -105.1145
11.38 | -105.1192
11.3838 | -105.1283
11.3876 | -105.1344
11.3914 | -105.1382
11.3951 | -105.1473
11.3989 | -105.1612
11.4027 | -105.1724
11.4065 | -105.1857
11.4102 | -105.1972
11.414 | -105.2107
11.4178 | -105.2233
11.4216 | -105.2359
11.4253 | -105.2516
11.4291 | -105.2649
11.4329 | -105.2809
11.4367 | -105.2917
11.4404 | -105.3052
11.4442 | -105.3175
11.448 | -105.3299
11.4518 | -105.3411
11.4555 | -105.3534
11.4593 | -105.3717
11.4631 | -105.3908
11.4669 | -105.4098
11.4706 | -105.4278
11.4744 | -105.4423
11.4782 | -105.4584
11.482 | -105.4727
11.4857 | -105.4834
11.4895 | -105.4921
11.4933 | -105.4995
11.4971 | -105.508
11.5008 | -105.5149
11.5046 | -105.5165
11.5084 | -105.5155
11.5121 | -105.5159
11.5159 | -105.5205
11.5197 | -105.5195
11.5235 | -105.51
11.5272 | -105.5021
11.531 | -105.4934
11.5348 | -105.4846
11.5386 | -105.4775
11.5423 | -105.4642
11.5461 | -105.4493
11.5499 | -105.4359
11.5537 | -105.4239
11.5574 | -105.4138
11.5612 | -105.4021
11.565 | -105.3907
11.5688 | -105.38
11.5725 | -105.3689
11.5763 | -105.3635
11.5801 | -105.3595
11.5839 | -105.3529
11.5876 | -105.3486
11.5914 | -105.3444
11.5952 | -105.3427
11.599 | -105.345
11.6027 | -105.3454
11.6065 | -105.3505
11.6103 | -105.3591
11.6141 | -105.3664
11.6178 | -105.3715
11.6216 | -105.3786
11.6254 | -105.3816
11.6292 | -105.3848
11.6329 | -105.3868
11.6367 | -105.3895
11.6405 | -105.3934
11.6443 | -105.3949
11.648 | -105.3948
11.6518 | -105.3922
11.6556 | -105.3882
11.6594 | -105.3842
11.6631 | -105.3799
11.6669 | -105.3699
11.6707 | -105.3614
11.6745 | -105.3558
11.6782 | -105.344
11.682 | -105.3281
11.6858 | -105.3121
11.6895 | -105.2995
11.6933 | -105.2871
11.6971 | -105.27
11.7009 | -105.2538
11.7046 | -105.2393
11.7084 | -105.2231
11.7122 | -105.2121
11.716 | -105.2027
11.7197 | -105.1919
11.7235 | -105.1837
11.7273 | -105.1769
11.7311 | -105.1726
11.7348 | -105.1696
11.7386 | -105.1722
11.7424 | -105.1774
11.7462 | -105.1837
11.7499 | -105.1891
11.7537 | -105.2026
11.7575 | -105.2156
11.7613 | -105.2246
11.765 | -105.2353
11.7688 | -105.2415
11.7726 | -105.251
11.7764 | -105.2571
11.7801 | -105.2597
11.7839 | -105.2623
11.7877 | -105.2636
11.7915 | -105.2588
11.7952 | -105.2539
11.799 | -105.247
11.8028 | -105.2366
11.8066 | -105.22
11.8103 | -105.1993
11.8141 | -105.1823
11.8179 | -105.1662
11.8217 | -105.149
11.8254 | -105.129
11.8292 | -105.1087
11.833 | -105.0914
11.8368 | -105.0803
11.8405 | -105.0687
11.8443 | -105.0601
11.8481 | -105.0604
11.8519 | -105.0617
11.8556 | -105.0609
11.8594 | -105.0646
11.8632 | -105.0752
11.8669 | -105.0929
11.8707 | -105.1126
11.8745 | -105.1353
11.8783 | -105.166
11.882 | -105.198
11.8858 | -105.2328
11.8896 | -105.2694
11.8934 | -105.3063
11.8971 | -105.3435
11.9009 | -105.3819
11.9047 | -105.419
11.9085 | -105.4524
11.9122 | -105.4846
11.916 | -105.5127
11.9198 | -105.5366
11.9236 | -105.5543
11.9273 | -105.5759
11.9311 | -105.5993
11.9349 | -105.6108
11.9387 | -105.6221
11.9424 | -105.6286
11.9462 | -105.6327
11.95 | -105.6393
11.9538 | -105.6441
11.9575 | -105.6494
11.9613 | -105.6536
11.9651 | -105.6587
11.9689 | -105.6606
11.9726 | -105.662
11.9764 | -105.6638
11.9802 | -105.6692
11.984 | -105.6752
11.9877 | -105.6748
11.9915 | -105.6709
11.9953 | -105.6646
11.9991 | -105.6575
12.0028 | -105.6497
12.0066 | -105.6383
12.0104 | -105.6246
12.0142 | -105.6154
12.0179 | -105.6031
12.0217 | -105.5867
12.0255 | -105.5754
12.0293 | -105.5655
12.033 | -105.5523
12.0368 | -105.5388
12.0406 | -105.5231
12.0443 | -105.5066
12.0481 | -105.4896
12.0519 | -105.4715
12.0557 | -105.4597
12.0594 | -105.4448
12.0632 | -105.4303
12.067 | -105.4164
12.0708 | -105.4044
12.0745 | -105.3899
12.0783 | -105.3813
12.0821 | -105.3767
12.0859 | -105.3728
12.0896 | -105.3707
12.0934 | -105.3713
12.0972 | -105.3701
12.101 | -105.3677
12.1047 | -105.3669
12.1085 | -105.3679
12.1123 | -105.3667
12.1161 | -105.3624
12.1198 | -105.3547
12.1236 | -105.3464
12.1274 | -105.3401
12.1312 | -105.3327
12.1349 | -105.327
12.1387 | -105.3212
12.1425 | -105.3155
12.1463 | -105.3122
12.15 | -105.3093
12.1538 | -105.3052
12.1576 | -105.3028
12.1614 | -105.3004
12.1651 | -105.2994
12.1689 | -105.3014
12.1727 | -105.3093
12.1765 | -105.3199
12.1802 | -105.3325
12.184 | -105.3485
12.1878 | -105.3663
12.1916 | -105.3815
12.1953 | -105.4
12.1991 | -105.4205
12.2029 | -105.4367
12.2067 | -105.4569
12.2104 | -105.4796
12.2142 | -105.5018
12.218 | -105.525
12.2218 | -105.544
12.2255 | -105.5634
12.2293 | -105.5823
12.2331 | -105.5975
12.2368 | -105.6156
12.2406 | -105.6331
12.2444 | -105.6498
12.2482 | -105.6638
12.2519 | -105.6776
12.2557 | -105.686
12.2595 | -105.6901
12.2633 | -105.6946
12.267 | -105.6996
12.2708 | -105.7068
12.2746 | -105.7173
12.2784 | -105.7244
12.2821 | -105.732
12.2859 | -105.7437
12.2897 | -105.7544
12.2935 | -105.7684
12.2972 | -105.7841
12.301 | -105.7926
12.3048 | -105.7985
12.3086 | -105.8091
12.3123 | -105.8153
12.3161 | -105.8156
12.3199 | -105.8177
12.3237 | -105.816
12.3274 | -105.8142
12.3312 | -105.81
12.335 | -105.7966
12.3388 | -105.7814
12.3425 | -105.7654
12.3463 | -105.7416
12.3501 | -105.7087
12.3539 | -105.6756
12.3576 | -105.6413
12.3614 | -105.6086
12.3652 | -105.5767
12.369 | -105.543
12.3727 | -105.5099
12.3765 | -105.4788
12.3803 | -105.4506
12.3841 | -105.4236
12.3878 | -105.3985
12.3916 | -105.3728
12.3954 | -105.3469
12.3992 | -105.3305
12.4029 | -105.3179
12.4067 | -105.3055
12.4105 | -105.2951
12.4142 | -105.2914
12.418 | -105.2943
12.4218 | -105.2979
12.4256 | -105.3027
12.4293 | -105.3107
12.4331 | -105.3211
12.4369 | -105.3347
12.4407 | -105.3519
12.4444 | -105.368
12.4482 | -105.3837
12.452 | -105.3983
12.4558 | -105.4115
12.4595 | -105.424
12.4633 | -105.4324
12.4671 | -105.442
12.4709 | -105.4467
12.4746 | -105.4455
12.4784 | -105.4469
12.4822 | -105.4455
12.486 | -105.4377
12.4897 | -105.4342
12.4935 | -105.4323
12.4973 | -105.4245
12.5011 | -105.4186
12.5048 | -105.4113
12.5086 | -105.4061
12.5124 | -105.4072
12.5162 | -105.4046
12.5199 | -105.4024
12.5237 | -105.4113
12.5275 | -105.4194
12.5313 | -105.4298
12.535 | -105.4387
12.5388 | -105.4468
12.5426 | -105.4527
12.5464 | -105.4609
12.5501 | -105.4757
12.5539 | -105.4921
12.5577 | -105.5079
12.5615 | -105.5235
12.5652 | -105.5386
12.569 | -105.5568
12.5728 | -105.5793
12.5766 | -105.6046
12.5803 | -105.6229
12.5841 | -105.6393
12.5879 | -105.6583
12.5916 | -105.6758
12.5954 | -105.687
12.5992 | -105.6942
12.603 | -105.7025
12.6067 | -105.7106
12.6105 | -105.7161
12.6143 | -105.7203
12.6181 | -105.7216
12.6218 | -105.7173
12.6256 | -105.7128
12.6294 | -105.7052
12.6332 | -105.6968
12.6369 | -105.6908
12.6407 | -105.6856
12.6445 | -105.6824
12.6483 | -105.6752
12.652 | -105.6707
12.6558 | -105.6706
12.6596 | -105.6715
12.6634 | -105.6739
12.6671 | -105.6786
12.6709 | -105.6805
12.6747 | -105.6827
12.6785 | -105.6913
12.6822 | -105.7015
12.686 | -105.7101
12.6898 | -105.718
12.6936 | -105.7262
12.6973 | -105.7383
12.7011 | -105.7474
12.7049 | -105.752
12.7087 | -105.7521
12.7124 | -105.7525
12.7162 | -105.7504
12.72 | -105.746
12.7238 | -105.7395
12.7275 | -105.731
12.7313 | -105.7208
12.7351 | -105.7055
12.7389 | -105.6884
12.7426 | -105.6704
12.7464 | -105.6518
12.7502 | -105.6302
12.754 | -105.6043
12.7577 | -105.5778
12.7615 | -105.55
12.7653 | -105.5247
12.769 | -105.5031
12.7728 | -105.4789
12.7766 | -105.4535
12.7804 | -105.4303
12.7841 | -105.4102
12.7879 | -105.3919
12.7917 | -105.3755
12.7955 | -105.3609
12.7992 | -105.3472
12.803 | -105.3376
12.8068 | -105.332
12.8106 | -105.3289
12.8143 | -105.3289
12.8181 | -105.3307
12.8219 | -105.3331
12.8257 | -105.3402
12.8294 | -105.3504
12.8332 | -105.3623
12.837 | -105.3725
12.8408 | -105.3843
12.8445 | -105.3981
12.8483 | -105.4116
12.8521 | -105.4279
12.8559 | -105.4483
12.8596 | -105.4671
12.8634 | -105.4846
12.8672 | -105.5019
12.871 | -105.5202
12.8747 | -105.5367
12.8785 | -105.552
12.8823 | -105.5659
12.8861 | -105.5792
12.8898 | -105.5942
12.8936 | -105.6074
12.8974 | -105.6133
12.9012 | -105.6194
12.9049 | -105.6222
12.9087 | -105.6261
12.9125 | -105.6294
12.9163 | -105.6292
12.92 | -105.6275
12.9238 | -105.627
12.9276 | -105.6251
12.9314 | -105.6225
12.9351 | -105.6238
12.9389 | -105.6241
12.9427 | -105.6227
12.9465 | -105.6222
12.9502 | -105.6213
12.954 | -105.6178
12.9578 | -105.6199
12.9615 | -105.6237
12.9653 | -105.6255
12.9691 | -105.6257
12.9729 | -105.6267
12.9766 | -105.6287
12.9804 | -105.6322
12.9842 | -105.638
12.988 | -105.6421
12.9917 | -105.6468
12.9955 | -105.6527
12.9993 | -105.6541
13.0031 | -105.6575
13.0068 | -105.6647
13.0106 | -105.6696
13.0144 | -105.673
13.0182 | -105.6808
13.0219 | -105.6896
13.0257 | -105.6966
13.0295 | -105.7008
13.0333 | -105.7079
13.037 | -105.7118
13.0408 | -105.7148
13.0446 | -105.7196
13.0484 | -105.7214
13.0521 | -105.7239
13.0559 | -105.7274
13.0597 | -105.7323
13.0635 | -105.7328
13.0672 | -105.735
13.071 | -105.7345
13.0748 | -105.7342
13.0786 | -105.7348
13.0823 | -105.7353
13.0861 | -105.7365
13.0899 | -105.7394
13.0937 | -105.7442
13.0974 | -105.7456
13.1012 | -105.7468
13.105 | -105.7481
13.1088 | -105.7503
13.1125 | -105.75
13.1163 | -105.7531
13.1201 | -105.7549
13.1238 | -105.7516
13.1276 | -105.7493
13.1314 | -105.7517
13.1352 | -105.7542
13.1389 | -105.7582
13.1427 | -105.7595
13.1465 | -105.7588
13.1503 | -105.7608
13.154 | -105.7638
13.1578 | -105.7648
13.1616 | -105.7638
13.1654 | -105.7623
13.1691 | -105.7622
13.1729 | -105.7603
13.1767 | -105.7588
13.1805 | -105.7563
13.1842 | -105.7555
13.188 | -105.7584
13.1918 | -105.7594
13.1956 | -105.7561
13.1993 | -105.7521
13.2031 | -105.7476
13.2069 | -105.7424
13.2107 | -105.7392
13.2144 | -105.738
13.2182 | -105.7372
13.222 | -105.7347
13.2258 | -105.731
13.2295 | -105.7285
13.2333 | -105.7236
13.2371 | -105.716
13.2409 | -105.7063
13.2446 | -105.6949
13.2484 | -105.6853
13.2522 | -105.6749
13.256 | -105.6599
13.2597 | -105.6412
13.2635 | -105.6257
13.2673 | -105.6108
13.2711 | -105.5933
13.2748 | -105.5767
13.2786 | -105.5597
13.2824 | -105.5387
13.2862 | -105.5171
13.2899 | -105.4975
13.2937 | -105.4785
13.2975 | -105.4615
13.3012 | -105.4446
13.305 | -105.4285
13.3088 | -105.4126
13.3126 | -105.4007
13.3163 | -105.3952
13.3201 | -105.3913
13.3239 | -105.3848
13.3277 | -105.3823
13.3314 | -105.3801
13.3352 | -105.3816
13.339 | -105.387
13.3428 | -105.3912
13.3465 | -105.3993
13.3503 | -105.4138
13.3541 | -105.4276
13.3579 | -105.4403
13.3616 | -105.4516
13.3654 | -105.4664
13.3692 | -105.4791
13.373 | -105.4907
13.3767 | -105.5004
13.3805 | -105.5073
13.3843 | -105.5116
13.3881 | -105.5165
13.3918 | -105.5176
13.3956 | -105.5188
13.3994 | -105.5187
13.4032 | -105.5147
13.4069 | -105.5051
13.4107 | -105.4941
13.4145 | -105.4849
13.4183 | -105.4758
13.422 | -105.4615
13.4258 | -105.448
13.4296 | -105.4382
13.4334 | -105.4307
13.4371 | -105.4243
13.4409 | -105.4193
13.4447 | -105.4164
13.4485 | -105.4143
13.4522 | -105.4137
13.456 | -105.4146
13.4598 | -105.4171
13.4636 | -105.4231
13.4673 | -105.4336
13.4711 | -105.446
13.4749 | -105.458
13.4787 | -105.4688
13.4824 | -105.4808
13.4862 | -105.4951
13.49 | -105.5077
13.4937 | -105.5181
13.4975 | -105.5274
13.5013 | -105.5356
13.5051 | -105.5387
13.5088 | -105.5366
13.5126 | -105.5338
13.5164 | -105.5299
13.5202 | -105.523
13.5239 | -105.5155
13.5277 | -105.5046
13.5315 | -105.4909
13.5353 | -105.4786
13.539 | -105.4623
13.5428 | -105.4446
13.5466 | -105.4307
13.5504 | -105.4175
13.5541 | -105.4024
13.5579 | -105.3863
13.5617 | -105.3714
13.5655 | -105.3577
13.5692 | -105.3448
13.573 | -105.3314
13.5768 | -105.3174
13.5806 | -105.3054
13.5843 | -105.2969
13.5881 | -105.2877
13.5919 | -105.278
13.5957 | -105.2708
13.5994 | -105.2674
13.6032 | -105.2626
13.607 | -105.2588
13.6108 | -105.2646
13.6145 | -105.2658
13.6183 | -105.271
13.6221 | -105.2792
13.6259 | -105.2872
13.6296 | -105.2968
13.6334 | -105.309
13.6372 | -105.3234
13.641 | -105.341
13.6447 | -105.3592
13.6485 | -105.379
13.6523 | -105.3989
13.6561 | -105.4204
13.6598 | -105.4447
13.6636 | -105.4669
13.6674 | -105.489
13.6711 | -105.5149
13.6749 | -105.5396
13.6787 | -105.5638
13.6825 | -105.5841
13.6862 | -105.6045
13.69 | -105.6237
13.6938 | -105.6412
13.6976 | -105.6603
13.7013 | -105.6793
13.7051 | -105.694
13.7089 | -105.7033
13.7127 | -105.7118
13.7164 | -105.7192
13.7202 | -105.7254
13.724 | -105.731
13.7278 | -105.7328
13.7315 | -105.7337
13.7353 | -105.7335
13.7391 | -105.7308
13.7429 | -105.7267
13.7466 | -105.7224
13.7504 | -105.7172
13.7542 | -105.7105
13.758 | -105.7042
13.7617 | -105.6968
13.7655 | -105.6867
13.7693 | -105.6808
13.7731 | -105.6785
13.7768 | -105.6732
13.7806 | -105.6677
13.7844 | -105.6621
13.7882 | -105.6588
13.7919 | -105.6583
13.7957 | -105.6591
13.7995 | -105.6623
13.8033 | -105.6624
13.807 | -105.6643
13.8108 | -105.6696
13.8146 | -105.6712
13.8184 | -105.6744
13.8221 | -105.6751
13.8259 | -105.6735
13.8297 | -105.6719
13.8335 | -105.6695
13.8372 | -105.6663
13.841 | -105.6626
13.8448 | -105.6623
13.8485 | -105.6597
13.8523 | -105.655
13.8561 | -105.6457
13.8599 | -105.6361
13.8636 | -105.6278
13.8674 | -105.619
13.8712 | -105.6074
13.875 | -105.5978
13.8787 | -105.5834
13.8825 | -105.5696
13.8863 | -105.5564
13.8901 | -105.5447
13.8938 | -105.5288
13.8976 | -105.5095
13.9014 | -105.4963
13.9052 | -105.4841
13.9089 | -105.4757
13.9127 | -105.467
13.9165 | -105.4573
13.9203 | -105.4474
13.924 | -105.4373
13.9278 | -105.4295
13.9316 | -105.4222
13.9354 | -105.4181
13.9391 | -105.4139
13.9429 | -105.4074
13.9467 | -105.4003
13.9505 | -105.3986
13.9542 | -105.3982
13.958 | -105.3972
13.9618 | -105.3935
13.9656 | -105.3915
13.9693 | -105.3887
13.9731 | -105.3859
13.9769 | -105.3821
13.9807 | -105.3768
13.9844 | -105.3739
13.9882 | -105.369
13.992 | -105.3603
13.9958 | -105.3482
13.9995 | -105.3362
14.0033 | -105.3258
14.0071 | -105.3113
14.0109 | -105.2974
14.0146 | -105.2832
14.0184 | -105.2664
14.0222 | -105.2492
14.0259 | -105.232
14.0297 | -105.2106
14.0335 | -105.1903
14.0373 | -105.1691
14.041 | -105.1481
14.0448 | -105.1293
14.0486 | -105.1132
14.0524 | -105.0992
14.0561 | -105.0888
14.0599 | -105.0792
14.0637 | -105.0714
14.0675 | -105.0687
14.0712 | -105.0727
14.075 | -105.0763
14.0788 | -105.0849
14.0826 | -105.0975
14.0863 | -105.1123
14.0901 | -105.1295
14.0939 | -105.1527
14.0977 | -105.1758
14.1014 | -105.1958
14.1052 | -105.2204
14.109 | -105.2426
14.1128 | -105.2653
14.1165 | -105.2897
14.1203 | -105.3133
14.1241 | -105.3421
14.1279 | -105.3722
14.1316 | -105.4016
14.1354 | -105.4281
14.1392 | -105.4551
14.143 | -105.4827
14.1467 | -105.5087
14.1505 | -105.5308
14.1543 | -105.5402
14.1581 | -105.5618
14.1618 | -105.5805
14.1656 | -105.5984
14.1694 | -105.6125
14.1732 | -105.6251
14.1769 | -105.6315
14.1807 | -105.6402
14.1845 | -105.6479
14.1883 | -105.6517
14.192 | -105.6536
14.1958 | -105.654
14.1996 | -105.6562
14.2034 | -105.6575
14.2071 | -105.6555
14.2109 | -105.6535
14.2147 | -105.6491
14.2184 | -105.6446
14.2222 | -105.6367
14.226 | -105.6304
14.2298 | -105.6258
14.2335 | -105.6226
14.2373 | -105.6198
14.2411 | -105.6206
14.2449 | -105.6217
14.2486 | -105.6174
14.2524 | -105.6152
14.2562 | -105.6138
14.26 | -105.614
14.2637 | -105.6142
14.2675 | -105.6165
14.2713 | -105.62
14.2751 | -105.6218
14.2788 | -105.6292
14.2826 | -105.6363
14.2864 | -105.64
14.2902 | -105.6441
14.2939 | -105.6508
14.2977 | -105.6588
14.3015 | -105.6682
14.3053 | -105.6797
14.309 | -105.6897
14.3128 | -105.7001
14.3166 | -105.7101
14.3204 | -105.717
14.3241 | -105.7265
14.3279 | -105.7331
14.3317 | -105.7335
14.3355 | -105.7319
14.3392 | -105.734
14.343 | -105.7348
14.3468 | -105.735
14.3506 | -105.7355
14.3543 | -105.732
14.3581 | -105.7307
14.3619 | -105.7285
14.3657 | -105.7219
14.3694 | -105.7171
14.3732 | -105.7169
14.377 | -105.7163
14.3808 | -105.7115
14.3845 | -105.7078
14.3883 | -105.7064
14.3921 | -105.7043
14.3958 | -105.7047
14.3996 | -105.7034
14.4034 | -105.703
14.4072 | -105.7021
14.4109 | -105.7021
14.4147 | -105.7043
14.4185 | -105.7088
14.4223 | -105.712
14.426 | -105.7163
14.4298 | -105.725
14.4336 | -105.7321
14.4374 | -105.7348
14.4411 | -105.739
14.4449 | -105.742
14.4487 | -105.7435
14.4525 | -105.7474
14.4562 | -105.7542
14.46 | -105.757
14.4638 | -105.7573
14.4676 | -105.7603
14.4713 | -105.7643
14.4751 | -105.7663
14.4789 | -105.7676
14.4827 | -105.7717
14.4864 | -105.7724
14.4902 | -105.7732
14.494 | -105.7764
14.4978 | -105.7807
14.5015 | -105.7834
14.5053 | -105.7841
14.5091 | -105.7871
14.5129 | -105.7904
14.5166 | -105.7926
14.5204 | -105.7955
14.5242 | -105.7972
14.528 | -105.8037
14.5317 | -105.8076
14.5355 | -105.8102
14.5393 | -105.8189
14.5431 | -105.8288
14.5468 | -105.8328
14.5506 | -105.8362
14.5544 | -105.8436
14.5582 | -105.8501
14.5619 | -105.8589
14.5657 | -105.8676
14.5695 | -105.8725
14.5732 | -105.8781
14.577 | -105.8852
14.5808 | -105.8877
14.5846 | -105.8911
14.5883 | -105.8994
14.5921 | -105.9061
14.5959 | -105.9091
14.5997 | -105.9117
14.6034 | -105.9153
14.6072 | -105.9191
14.611 | -105.9194
14.6148 | -105.9183
14.6185 | -105.9139
14.6223 | -105.9099
14.6261 | -105.9012
14.6299 | -105.8968
14.6336 | -105.8904
14.6374 | -105.8811
14.6412 | -105.872
14.645 | -105.8653
14.6487 | -105.8551
14.6525 | -105.8418
14.6563 | -105.8282
14.6601 | -105.8123
14.6638 | -105.7931
14.6676 | -105.7789
14.6714 | -105.7689
14.6752 | -105.7544
14.6789 | -105.7392
14.6827 | -105.725
14.6865 | -105.7097
14.6903 | -105.6932
14.694 | -105.6775
14.6978 | -105.6655
14.7016 | -105.6651
14.7054 | -105.6553
14.7091 | -105.6439
14.7129 | -105.6333
14.7167 | -105.6287
14.7205 | -105.6241
14.7242 | -105.6188
14.728 | -105.6184
14.7318 | -105.6186
14.7356 | -105.6193
14.7393 | -105.6186
14.7431 | -105.6177
14.7469 | -105.6187
14.7506 | -105.621
14.7544 | -105.6248
14.7582 | -105.6292
14.762 | -105.6316
14.7657 | -105.6327
14.7695 | -105.6344
14.7733 | -105.6344
14.7771 | -105.6358
14.7808 | -105.6383
14.7846 | -105.639
14.7884 | -105.6421
14.7922 | -105.6429
14.7959 | -105.6425
14.7997 | -105.6435
14.8035 | -105.6451
14.8073 | -105.6464
14.811 | -105.652
14.8148 | -105.6546
14.8186 | -105.6592
14.8224 | -105.6645
14.8261 | -105.669
14.8299 | -105.6776
14.8337 | -105.6841
14.8375 | -105.6911
14.8412 | -105.696
14.845 | -105.7038
14.8488 | -105.7142
14.8526 | -105.723
14.8563 | -105.7315
14.8601 | -105.7406
14.8639 | -105.7505
14.8677 | -105.7602
14.8714 | -105.7699
14.8752 | -105.7796
14.879 | -105.7847
14.8828 | -105.7893
14.8865 | -105.7921
14.8903 | -105.7923
14.8941 | -105.7933
14.8979 | -105.7929
14.9016 | -105.7915
14.9054 | -105.7907
14.9092 | -105.7863
14.913 | -105.7812
14.9167 | -105.7764
14.9205 | -105.7742
14.9243 | -105.7717
14.928 | -105.7659
14.9318 | -105.757
14.9356 | -105.749
14.9394 | -105.7389
14.9431 | -105.7253
14.9469 | -105.7145
14.9507 | -105.7035
14.9545 | -105.6929
14.9582 | -105.6826
14.962 | -105.6692
14.9658 | -105.6582
14.9696 | -105.6495
14.9733 | -105.6431
14.9771 | -105.6393
14.9809 | -105.6342
14.9847 | -105.63
14.9884 | -105.6272
14.9922 | -105.6219
14.996 | -105.6186
14.9998 | -105.6174
15.0035 | -105.6158
15.0073 | -105.617
15.0111 | -105.6168
15.0149 | -105.6138
15.0186 | -105.6125
15.0224 | -105.6131
15.0262 | -105.6127
15.03 | -105.6143
15.0337 | -105.6176
15.0375 | -105.6244
15.0413 | -105.6293
15.0451 | -105.632
15.0488 | -105.6387
15.0526 | -105.6465
15.0564 | -105.6554
15.0602 | -105.6648
15.0639 | -105.6738
15.0677 | -105.6861
15.0715 | -105.6985
15.0753 | -105.7112
15.079 | -105.7257
15.0828 | -105.7408
15.0866 | -105.7566
15.0904 | -105.7757
15.0941 | -105.7945
15.0979 | -105.8113
15.1017 | -105.8308
15.1054 | -105.8536
15.1092 | -105.875
15.113 | -105.8971
15.1168 | -105.9168
15.1205 | -105.9364
15.1243 | -105.955
15.1281 | -105.9725
15.1319 | -105.9866
15.1356 | -106.0031
15.1394 | -106.0196
15.1432 | -106.0345
15.147 | -106.0475
15.1507 | -106.0607
15.1545 | -106.0757
15.1583 | -106.0903
15.1621 | -106.099
15.1658 | -106.1078
15.1696 | -106.1152
15.1734 | -106.1207
15.1772 | -106.1282
15.1809 | -106.1276
15.1847 | -106.125
15.1885 | -106.1245
15.1923 | -106.1223
15.196 | -106.1164
15.1998 | -106.1095
15.2036 | -106.1009
15.2074 | -106.0906
15.2111 | -106.0791
15.2149 | -106.0646
15.2187 | -106.0477
15.2225 | -106.0295
15.2262 | -106.0116
15.23 | -105.9936
15.2338 | -105.9724
15.2376 | -105.9476
15.2413 | -105.9231
15.2451 | -105.8942
15.2489 | -105.8707
15.2527 | -105.8446
15.2564 | -105.82
15.2602 | -105.7999
15.264 | -105.7822
15.2678 | -105.7635
15.2715 | -105.7437
15.2753 | -105.7251
15.2791 | -105.7063
15.2828 | -105.6892
15.2866 | -105.6707
15.2904 | -105.6537
15.2942 | -105.6418
15.2979 | -105.6313
15.3017 | -105.6204
15.3055 | -105.6115
15.3093 | -105.6022
15.313 | -105.594
15.3168 | -105.5862
15.3206 | -105.5782
15.3244 | -105.5705
15.3281 | -105.5619
15.3319 | -105.5544
15.3357 | -105.5493
15.3395 | -105.5467
15.3432 | -105.5445
15.347 | -105.5432
15.3508 | -105.5421
15.3546 | -105.5397
15.3583 | -105.5394
15.3621 | -105.5431
15.3659 | -105.5478
15.3697 | -105.5502
15.3734 | -105.5513
15.3772 | -105.5529
15.381 | -105.556
15.3848 | -105.5595
15.3885 | -105.5635
15.3923 | -105.57
15.3961 | -105.5759
15.3999 | -105.5834
15.4036 | -105.5917
15.4074 | -105.5977
15.4112 | -105.6036
15.415 | -105.6123
15.4187 | -105.6204
15.4225 | -105.6302
15.4263 | -105.6413
15.4301 | -105.6517
15.4338 | -105.6619
15.4376 | -105.669
15.4414 | -105.6787
15.4452 | -105.6912
15.4489 | -105.703
15.4527 | -105.714
15.4565 | -105.725
15.4603 | -105.7367
15.464 | -105.7519
15.4678 | -105.766
15.4716 | -105.7754
15.4753 | -105.7849
15.4791 | -105.7942
15.4829 | -105.8029
15.4867 | -105.8077
15.4904 | -105.8111
15.4942 | -105.8165
15.498 | -105.8206
15.5018 | -105.8244
15.5055 | -105.8237
15.5093 | -105.8238
15.5131 | -105.8223
15.5169 | -105.817
15.5206 | -105.8111
15.5244 | -105.8062
15.5282 | -105.7977
15.532 | -105.787
15.5357 | -105.7761
15.5395 | -105.7648
15.5433 | -105.7539
15.5471 | -105.7419
15.5508 | -105.7292
15.5546 | -105.7129
15.5584 | -105.6953
15.5622 | -105.6754
15.5659 | -105.6533
15.5697 | -105.6299
15.5735 | -105.6085
15.5773 | -105.5872
15.581 | -105.5666
15.5848 | -105.545
15.5886 | -105.5219
15.5924 | -105.5002
15.5961 | -105.4792
15.5999 | -105.4578
15.6037 | -105.4328
15.6075 | -105.4085
15.6112 | -105.3854
15.615 | -105.3628
15.6188 | -105.3409
15.6226 | -105.318
15.6263 | -105.2958
15.6301 | -105.2751
15.6339 | -105.2554
15.6377 | -105.2379
15.6414 | -105.2218
15.6452 | -105.2049
15.649 | -105.1888
15.6527 | -105.1769
15.6565 | -105.1665
15.6603 | -105.1576
15.6641 | -105.1481
15.6678 | -105.1386
15.6716 | -105.1311
15.6754 | -105.1275
15.6792 | -105.1236
15.6829 | -105.1198
15.6867 | -105.1168
15.6905 | -105.1143
15.6943 | -105.1127
15.698 | -105.1153
15.7018 | -105.1177
15.7056 | -105.1191
15.7094 | -105.1236
15.7131 | -105.1295
15.7169 | -105.1348
15.7207 | -105.1432
15.7245 | -105.1499
15.7282 | -105.1584
15.732 | -105.1658
15.7358 | -105.175
15.7396 | -105.184
15.7433 | -105.1896
15.7471 | -105.1972
15.7509 | -105.2084
15.7547 | -105.2188
15.7584 | -105.2254
15.7622 | -105.2328
15.766 | -105.2406
15.7698 | -105.2442
15.7735 | -105.2484
15.7773 | -105.2549
15.7811 | -105.2578
15.7849 | -105.2581
15.7886 | -105.2543
15.7924 | -105.2556
15.7962 | -105.256
15.8 | -105.253
15.8037 | -105.2501
15.8075 | -105.244
15.8113 | -105.2352
15.8151 | -105.2261
15.8188 | -105.2174
15.8226 | -105.2103
15.8264 | -105.2023
15.8301 | -105.1981
15.8339 | -105.192
15.8377 | -105.1857
15.8415 | -105.1794
15.8452 | -105.1722
15.849 | -105.168
15.8528 | -105.1634
15.8566 | -105.159
15.8603 | -105.1575
15.8641 | -105.1551
15.8679 | -105.1527
15.8717 | -105.1526
15.8754 | -105.1534
15.8792 | -105.1544
15.883 | -105.158
15.8868 | -105.1641
15.8905 | -105.1714
15.8943 | -105.1773
15.8981 | -105.1838
15.9019 | -105.1935
15.9056 | -105.2048
15.9094 | -105.2153
15.9132 | -105.2287
15.917 | -105.2453
15.9207 | -105.2642
15.9245 | -105.2827
15.9283 | -105.3002
15.9321 | -105.3205
15.9358 | -105.3424
15.9396 | -105.3636
15.9434 | -105.3843
15.9472 | -105.403
15.9509 | -105.4229
15.9547 | -105.4416
15.9585 | -105.4608
15.9623 | -105.4771
15.966 | -105.4918
15.9698 | -105.5084
15.9736 | -105.5211
15.9774 | -105.5319
15.9811 | -105.5454
15.9849 | -105.5587
15.9887 | -105.569
15.9925 | -105.5775
15.9962 | -105.586
16 | -105.5909
16.0038 | -105.5915
16.0075 | -105.5922
16.0113 | -105.5951
16.0151 | -105.5958
16.0189 | -105.5954
16.0226 | -105.5963
16.0264 | -105.596
16.0302 | -105.5958
16.034 | -105.5935
16.0377 | -105.5933
16.0415 | -105.5948
16.0453 | -105.5992
16.0491 | -105.6011
16.0528 | -105.5999
16.0566 | -105.6004
16.0604 | -105.5996
16.0642 | -105.6003
16.0679 | -105.6024
16.0717 | -105.6065
16.0755 | -105.6115
16.0793 | -105.6152
16.083 | -105.619
16.0868 | -105.6237
16.0906 | -105.6278
16.0944 | -105.6322
16.0981 | -105.6365
16.1019 | -105.6417
16.1057 | -105.6475
16.1095 | -105.6507
16.1132 | -105.6531
16.117 | -105.6553
16.1208 | -105.6574
16.1246 | -105.6614
16.1283 | -105.6664
16.1321 | -105.6687
16.1359 | -105.6702
16.1397 | -105.6727
16.1434 | -105.6716
16.1472 | -105.669
16.151 | -105.6686
16.1548 | -105.6671
16.1585 | -105.6676
16.1623 | -105.6662
16.1661 | -105.6624
16.1699 | -105.662
16.1736 | -105.6648
16.1774 | -105.6656
16.1812 | -105.6658
16.1849 | -105.6687
16.1887 | -105.669
16.1925 | -105.674
16.1963 | -105.6784
16.2 | -105.68
16.2038 | -105.6806
16.2076 | -105.6828
16.2114 | -105.6809
16.2151 | -105.6814
16.2189 | -105.6824
16.2227 | -105.6812
16.2265 | -105.6786
16.2302 | -105.6786
16.234 | -105.6784
16.2378 | -105.6764
16.2416 | -105.67
16.2453 | -105.6632
16.2491 | -105.6543
16.2529 | -105.6444
16.2567 | -105.6369
16.2604 | -105.6287
16.2642 | -105.6207
16.268 | -105.6105
16.2718 | -105.5995
16.2755 | -105.585
16.2793 | -105.5722
16.2831 | -105.5628
16.2869 | -105.5514
16.2906 | -105.5417
16.2944 | -105.5295
16.2982 | -105.5158
16.302 | -105.5046
16.3057 | -105.496
16.3095 | -105.4874
16.3133 | -105.479
16.3171 | -105.4714
16.3208 | -105.4627
16.3246 | -105.4565
16.3284 | -105.4515
16.3322 | -105.445
16.3359 | -105.4455
16.3397 | -105.4405
16.3435 | -105.4376
16.3473 | -105.435
16.351 | -105.4341
16.3548 | -105.4313
16.3586 | -105.4274
16.3623 | -105.4244
16.3661 | -105.4235
16.3699 | -105.4246
16.3737 | -105.4241
16.3774 | -105.4214
16.3812 | -105.4193
16.385 | -105.4168
16.3888 | -105.4148
16.3925 | -105.4141
16.3963 | -105.4123
16.4001 | -105.4117
16.4039 | -105.4099
16.4076 | -105.4076
16.4114 | -105.4071
16.4152 | -105.4088
16.419 | -105.4093
16.4227 | -105.4123
16.4265 | -105.4127
16.4303 | -105.4154
16.4341 | -105.4189
16.4378 | -105.4257
16.4416 | -105.4315
16.4454 | -105.439
16.4492 | -105.4473
16.4529 | -105.4554
16.4567 | -105.4668
16.4605 | -105.479
16.4643 | -105.489
16.468 | -105.5028
16.4718 | -105.5179
16.4756 | -105.5323
16.4794 | -105.5452
16.4831 | -105.5604
16.4869 | -105.5749
16.4907 | -105.5867
16.4945 | -105.5989
16.4982 | -105.6123
16.502 | -105.6258
16.5058 | -105.6392
16.5096 | -105.6531
16.5133 | -105.6634
16.5171 | -105.6698
16.5209 | -105.6788
16.5247 | -105.6854
16.5284 | -105.6938
16.5322 | -105.7008
16.536 | -105.7054
16.5398 | -105.71
16.5435 | -105.7165
16.5473 | -105.7232
16.5511 | -105.7292
16.5548 | -105.7344
16.5586 | -105.7346
16.5624 | -105.7367
16.5662 | -105.7418
16.5699 | -105.7451
16.5737 | -105.7497
16.5775 | -105.7559
16.5813 | -105.7625
16.585 | -105.7707
16.5888 | -105.7794
16.5926 | -105.7916
16.5964 | -105.8043
16.6001 | -105.8156
16.6039 | -105.8276
16.6077 | -105.8415
16.6115 | -105.8557
16.6152 | -105.871
16.619 | -105.8851
16.6228 | -105.8987
16.6266 | -105.9142
16.6303 | -105.9306
16.6341 | -105.9497
16.6379 | -105.9693
16.6417 | -105.988
16.6454 | -106.0062
16.6492 | -106.0231
16.653 | -106.0384
16.6568 | -106.0541
16.6605 | -106.0653
16.6643 | -106.0773
16.6681 | -106.0893
16.6719 | -106.1038
16.6756 | -106.1155
16.6794 | -106.1259
16.6832 | -106.1339
16.687 | -106.1394
16.6907 | -106.1445
16.6945 | -106.1467
16.6983 | -106.147
16.7021 | -106.1456
16.7058 | -106.1405
16.7096 | -106.1349
16.7134 | -106.1305
16.7172 | -106.1245
16.7209 | -106.1171
16.7247 | -106.1096
16.7285 | -106.1017
16.7322 | -106.0938
16.736 | -106.0832
16.7398 | -106.0713
16.7436 | -106.0598
16.7473 | -106.0482
16.7511 | -106.0384
16.7549 | -106.0281
16.7587 | -106.0148
16.7624 | -106.0048
16.7662 | -105.9958
16.77 | -105.9874
16.7738 | -105.9794
16.7775 | -105.9706
16.7813 | -105.9621
16.7851 | -105.9557
16.7889 | -105.9511
16.7926 | -105.9478
16.7964 | -105.9464
16.8002 | -105.9452
16.804 | -105.9437
16.8077 | -105.9436
16.8115 | -105.9413
16.8153 | -105.9366
16.8191 | -105.9344
16.8228 | -105.9341
16.8266 | -105.9337
16.8304 | -105.9333
16.8342 | -105.931
16.8379 | -105.9306
16.8417 | -105.9307
16.8455 | -105.9311
16.8493 | -105.9336
16.853 | -105.9363
16.8568 | -105.9354
16.8606 | -105.934
16.8644 | -105.9314
16.8681 | -105.9281
16.8719 | -105.9259
16.8757 | -105.9208
16.8795 | -105.9156
16.8832 | -105.912
16.887 | -105.9046
16.8908 | -105.8997
16.8946 | -105.8964
16.8983 | -105.8915
16.9021 | -105.8899
16.9059 | -105.8842
16.9096 | -105.877
16.9134 | -105.8694
16.9172 | -105.8648
16.921 | -105.8587
16.9247 | -105.8537
16.9285 | -105.8475
16.9323 | -105.8426
16.9361 | -105.837
16.9398 | -105.8321
16.9436 | -105.8295
16.9474 | -105.8258
16.9512 | -105.8212
16.9549 | -105.8166
16.9587 | -105.8114
16.9625 | -105.8062
16.9663 | -105.7983
16.97 | -105.7877
16.9738 | -105.7794
16.9776 | -105.7707
16.9814 | -105.765
16.9851 | -105.7597
16.9889 | -105.7524
16.9927 | -105.7468
16.9965 | -105.7434
17.0002 | -105.7385
17.004 | -105.7322
17.0078 | -105.7268
17.0116 | -105.7227
17.0153 | -105.7196
17.0191 | -105.7182
17.0229 | -105.7165
17.0267 | -105.7146
17.0304 | -105.7119
17.0342 | -105.7106
17.038 | -105.7112
17.0418 | -105.7104
17.0455 | -105.7087
17.0493 | -105.7075
17.0531 | -105.7078
17.0569 | -105.7109
17.0606 | -105.7109
17.0644 | -105.7092
17.0682 | -105.7083
17.072 | -105.7076
17.0757 | -105.7084
17.0795 | -105.708
17.0833 | -105.7067
17.087 | -105.7051
17.0908 | -105.7053
17.0946 | -105.7057
17.0984 | -105.7076
17.1021 | -105.7087
17.1059 | -105.7096
17.1097 | -105.7106
17.1135 | -105.7124
17.1172 | -105.7177
17.121 | -105.7234
17.1248 | -105.7283
17.1286 | -105.732
17.1323 | -105.7394
17.1361 | -105.7454
17.1399 | -105.7505
17.1437 | -105.7569
17.1474 | -105.7629
17.1512 | -105.7702
17.155 | -105.7791
17.1588 | -105.7878
17.1625 | -105.7923
17.1663 | -105.7984
17.1701 | -105.8041
17.1739 | -105.8087
17.1776 | -105.8125
17.1814 | -105.8189
17.1852 | -105.8241
17.189 | -105.8261
17.1927 | -105.8302
17.1965 | -105.8314
17.2003 | -105.8338
17.2041 | -105.8384
17.2078 | -105.8423
17.2116 | -105.8468
17.2154 | -105.8555
17.2192 | -105.8589
17.2229 | -105.8613
17.2267 | -105.8634
17.2305 | -105.8643
17.2343 | -105.8669
17.238 | -105.8668
17.2418 | -105.865
17.2456 | -105.8639
17.2494 | -105.8633
17.2531 | -105.8602
17.2569 | -105.8587
17.2607 | -105.857
17.2644 | -105.8541
17.2682 | -105.85
17.272 | -105.8442
17.2758 | -105.8397
17.2795 | -105.8364
17.2833 | -105.8335
17.2871 | -105.8317
17.2909 | -105.829
17.2946 | -105.8275
17.2984 | -105.8254
17.3022 | -105.8243
17.306 | -105.8247
17.3097 | -105.8247
17.3135 | -105.823
17.3173 | -105.8186
17.3211 | -105.8145
17.3248 | -105.813
17.3286 | -105.814
17.3324 | -105.8143
17.3362 | -105.8131
17.3399 | -105.8137
17.3437 | -105.8163
17.3475 | -105.8168
17.3513 | -105.8163
17.355 | -105.8169
17.3588 | -105.8149
17.3626 | -105.812
17.3664 | -105.8094
17.3701 | -105.8076
17.3739 | -105.8072
17.3777 | -105.8068
17.3815 | -105.805
17.3852 | -105.8028
17.389 | -105.8021
17.3928 | -105.7999
17.3966 | -105.7972
17.4003 | -105.7926
17.4041 | -105.7871
17.4079 | -105.7857
17.4117 | -105.7852
17.4154 | -105.7855
17.4192 | -105.7854
17.423 | -105.7812
17.4268 | -105.7772
17.4305 | -105.773
17.4343 | -105.7701
17.4381 | -105.7687
17.4418 | -105.762
17.4456 | -105.7583
17.4494 | -105.7539
17.4532 | -105.7485
17.4569 | -105.7429
17.4607 | -105.7381
17.4645 | -105.7311
17.4683 | -105.7285
17.472 | -105.7255
17.4758 | -105.7197
17.4796 | -105.715
17.4834 | -105.7112
17.4871 | -105.708
17.4909 | -105.705
17.4947 | -105.7018
17.4985 | -105.6986
17.5022 | -105.6991
17.506 | -105.6975
17.5098 | -105.6945
17.5136 | -105.694
17.5173 | -105.6945
17.5211 | -105.6982
17.5249 | -105.702
17.5287 | -105.7042
17.5324 | -105.7059
17.5362 | -105.7074
17.54 | -105.7098
17.5438 | -105.7127
17.5475 | -105.7187
17.5513 | -105.7238
17.5551 | -105.7267
17.5589 | -105.7302
17.5626 | -105.7358
17.5664 | -105.7442
17.5702 | -105.7545
17.574 | -105.7617
17.5777 | -105.7661
17.5815 | -105.7712
17.5853 | -105.7759
17.5891 | -105.7795
17.5928 | -105.7815
17.5966 | -105.7854
17.6004 | -105.7883
17.6042 | -105.789
17.6079 | -105.7913
17.6117 | -105.7939
17.6155 | -105.7951
17.6193 | -105.7924
17.623 | -105.7902
17.6268 | -105.7856
17.6306 | -105.7795
17.6343 | -105.772
17.6381 | -105.7617
17.6419 | -105.7506
17.6457 | -105.7407
17.6494 | -105.7336
17.6532 | -105.7266
17.657 | -105.7145
17.6608 | -105.7004
17.6645 | -105.6865
17.6683 | -105.6734
17.6721 | -105.6622
17.6759 | -105.6529
17.6796 | -105.6458
17.6834 | -105.6361
17.6872 | -105.6266
17.691 | -105.6166
17.6947 | -105.6092
17.6985 | -105.6021
17.7023 | -105.5947
17.7061 | -105.5881
17.7098 | -105.5828
17.7136 | -105.5806
17.7174 | -105.5806
17.7212 | -105.5781
17.7249 | -105.5771
17.7287 | -105.5794
17.7325 | -105.583
17.7363 | -105.5878
17.74 | -105.5925
17.7438 | -105.598
17.7476 | -105.6052
17.7514 | -105.6112
17.7551 | -105.6143
17.7589 | -105.6182
17.7627 | -105.627
17.7665 | -105.6332
17.7702 | -105.641
17.774 | -105.6476
17.7778 | -105.6545
17.7816 | -105.6617
17.7853 | -105.6698
17.7891 | -105.6773
17.7929 | -105.6835
17.7966 | -105.6901
17.8004 | -105.6947
17.8042 | -105.699
17.808 | -105.708
17.8117 | -105.7174
17.8155 | -105.7258
17.8193 | -105.7352
17.8231 | -105.7428
17.8268 | -105.7495
17.8306 | -105.7548
17.8344 | -105.7606
17.8382 | -105.7684
17.8419 | -105.7759
17.8457 | -105.784
17.8495 | -105.7877
17.8533 | -105.7925
17.857 | -105.7975
17.8608 | -105.8018
17.8646 | -105.804
17.8684 | -105.8043
17.8721 | -105.8068
17.8759 | -105.8081
17.8797 | -105.807
17.8835 | -105.805
17.8872 | -105.8041
17.891 | -105.8048
17.8948 | -105.8029
17.8986 | -105.7963
17.9023 | -105.7877
17.9061 | -105.7773
17.9099 | -105.7689
17.9137 | -105.7608
17.9174 | -105.7485
17.9212 | -105.7373
17.925 | -105.7255
17.9288 | -105.7123
17.9325 | -105.6973
17.9363 | -105.6822
17.9401 | -105.6679
17.9439 | -105.6519
17.9476 | -105.6354
17.9514 | -105.6187
17.9552 | -105.6012
17.959 | -105.5837
17.9627 | -105.5646
17.9665 | -105.5466
17.9703 | -105.5283
17.974 | -105.5093
17.9778 | -105.4913
17.9816 | -105.4755
17.9854 | -105.4594
17.9891 | -105.4436
17.9929 | -105.4274
17.9967 | -105.4145
18.0005 | -105.4039
18.0042 | -105.3949
18.008 | -105.3876
18.0118 | -105.3806
18.0156 | -105.377
18.0193 | -105.3736
18.0231 | -105.3727
18.0269 | -105.371
18.0307 | -105.3705
18.0344 | -105.3712
18.0382 | -105.3743
18.042 | -105.3804
18.0458 | -105.3895
18.0495 | -105.4006
18.0533 | -105.4118
18.0571 | -105.4259
18.0609 | -105.4414
18.0646 | -105.4555
18.0684 | -105.4687
18.0722 | -105.4843
18.076 | -105.4994
18.0797 | -105.5161
18.0835 | -105.534
18.0873 | -105.5537
18.0911 | -105.5734
18.0948 | -105.592
18.0986 | -105.6091
18.1024 | -105.6281
18.1062 | -105.6504
18.1099 | -105.6708
18.1137 | -105.6882
18.1175 | -105.7035
18.1213 | -105.7191
18.125 | -105.737
18.1288 | -105.7528
18.1326 | -105.7641
18.1364 | -105.7756
18.1401 | -105.7852
18.1439 | -105.7934
18.1477 | -105.7971
18.1515 | -105.8018
18.1552 | -105.8107
18.159 | -105.8162
18.1628 | -105.819
18.1665 | -105.8198
18.1703 | -105.8184
18.1741 | -105.8167
18.1779 | -105.8148
18.1816 | -105.8102
18.1854 | -105.8044
18.1892 | -105.8017
18.193 | -105.7997
18.1967 | -105.7956
18.2005 | -105.7924
18.2043 | -105.7884
18.2081 | -105.7841
18.2118 | -105.7806
18.2156 | -105.7779
18.2194 | -105.7754
18.2232 | -105.7691
18.2269 | -105.7629
18.2307 | -105.7582
18.2345 | -105.7535
18.2383 | -105.751
18.242 | -105.7455
18.2458 | -105.7397
18.2496 | -105.7325
18.2534 | -105.728
18.2571 | -105.7236
18.2609 | -105.7173
18.2647 | -105.7101
18.2685 | -105.7041
18.2722 | -105.6991
18.276 | -105.6907
18.2798 | -105.6798
18.2836 | -105.6694
18.2873 | -105.662
18.2911 | -105.6548
18.2949 | -105.646
18.2987 | -105.6363
18.3024 | -105.6245
18.3062 | -105.6119
18.31 | -105.6004
18.3138 | -105.5886
18.3175 | -105.5756
18.3213 | -105.5612
18.3251 | -105.5484
18.3289 | -105.5364
18.3326 | -105.522
18.3364 | -105.5068
18.3402 | -105.4924
18.3439 | -105.4767
18.3477 | -105.4626
18.3515 | -105.4491
18.3553 | -105.4354
18.359 | -105.4228
18.3628 | -105.4073
18.3666 | -105.3901
18.3704 | -105.3752
18.3741 | -105.3638
18.3779 | -105.3533
18.3817 | -105.3402
18.3855 | -105.3255
18.3892 | -105.3138
18.393 | -105.3052
18.3968 | -105.2994
18.4006 | -105.2946
18.4043 | -105.2878
18.4081 | -105.2839
18.4119 | -105.2798
18.4157 | -105.2743
18.4194 | -105.2679
18.4232 | -105.2636
18.427 | -105.2609
18.4308 | -105.2606
18.4345 | -105.2625
18.4383 | -105.2596
18.4421 | -105.2578
18.4459 | -105.2551
18.4496 | -105.2498
18.4534 | -105.2461
18.4572 | -105.243
18.461 | -105.2376
18.4647 | -105.2355
18.4685 | -105.2335
18.4723 | -105.2274
18.4761 | -105.2189
18.4798 | -105.2124
18.4836 | -105.2086
18.4874 | -105.2027
18.4912 | -105.1949
18.4949 | -105.1889
18.4987 | -105.1816
18.5025 | -105.176
18.5063 | -105.171
18.51 | -105.1681
18.5138 | -105.1656
18.5176 | -105.1607
18.5213 | -105.1545
18.5251 | -105.1478
18.5289 | -105.1428
18.5327 | -105.1394
18.5364 | -105.1357
18.5402 | -105.1338
18.544 | -105.1315
18.5478 | -105.1279
18.5515 | -105.122
18.5553 | -105.1174
18.5591 | -105.1111
18.5629 | -105.1052
18.5666 | -105.1011
18.5704 | -105.097
18.5742 | -105.0928
18.578 | -105.0882
18.5817 | -105.0818
18.5855 | -105.0757
18.5893 | -105.0682
18.5931 | -105.0615
18.5968 | -105.0558
18.6006 | -105.0494
18.6044 | -105.0425
18.6082 | -105.0358
18.6119 | -105.0283
18.6157 | -105.0208
18.6195 | -105.0114
18.6233 | -105.0034
18.627 | -104.997
18.6308 | -104.9908
18.6346 | -104.9848
18.6384 | -104.9789
18.6421 | -104.9745
18.6459 | -104.9726
18.6497 | -104.9706
18.6535 | -104.968
18.6572 | -104.9653
18.661 | -104.9631
18.6648 | -104.9635
18.6686 | -104.966
18.6723 | -104.9677
18.6761 | -104.9712
18.6799 | -104.9769
18.6837 | -104.985
18.6874 | -104.9948
18.6912 | -105.0046
18.695 | -105.0138
18.6988 | -105.0219
18.7025 | -105.0338
18.7063 | -105.0471
18.7101 | -105.0628
18.7138 | -105.0794
18.7176 | -105.0949
18.7214 | -105.1102
18.7252 | -105.129
18.7289 | -105.1461
18.7327 | -105.1625
18.7365 | -105.1812
18.7403 | -105.2006
18.744 | -105.2164
18.7478 | -105.229
18.7516 | -105.2431
18.7554 | -105.2565
18.7591 | -105.2665
18.7629 | -105.2757
18.7667 | -105.2828
18.7705 | -105.2889
18.7742 | -105.2918
18.778 | -105.2929
18.7818 | -105.2944
18.7856 | -105.2947
18.7893 | -105.2941
18.7931 | -105.293
18.7969 | -105.2914
18.8007 | -105.2865
18.8044 | -105.2825
18.8082 | -105.2783
18.812 | -105.275
18.8158 | -105.271
18.8195 | -105.2662
18.8233 | -105.2608
18.8271 | -105.2553
18.8309 | -105.2478
18.8346 | -105.2396
18.8384 | -105.2333
18.8422 | -105.2275
18.846 | -105.2224
18.8497 | -105.2199
18.8535 | -105.2186
18.8573 | -105.2201
18.8611 | -105.2211
18.8648 | -105.2211
18.8686 | -105.2235
18.8724 | -105.2258
18.8761 | -105.2287
18.8799 | -105.2323
18.8837 | -105.2341
18.8875 | -105.2389
18.8912 | -105.245
18.895 | -105.2523
18.8988 | -105.2602
18.9026 | -105.2678
18.9063 | -105.2735
18.9101 | -105.276
18.9139 | -105.2809
18.9177 | -105.2865
18.9214 | -105.2914
18.9252 | -105.2958
18.929 | -105.3006
18.9328 | -105.3036
18.9365 | -105.3039
18.9403 | -105.3044
18.9441 | -105.3048
18.9479 | -105.3019
18.9516 | -105.2989
18.9554 | -105.2955
18.9592 | -105.2926
18.963 | -105.2885
18.9667 | -105.2821
18.9705 | -105.2758
18.9743 | -105.2682
18.9781 | -105.2581
18.9818 | -105.2477
18.9856 | -105.2379
18.9894 | -105.2298
18.9932 | -105.2196
18.9969 | -105.2078
19.0007 | -105.1958
19.0045 | -105.1868
19.0083 | -105.178
19.012 | -105.1707
19.0158 | -105.1654
19.0196 | -105.1585
19.0234 | -105.1513
19.0271 | -105.1472
19.0309 | -105.1432
19.0347 | -105.139
19.0385 | -105.1375
19.0422 | -105.137
19.046 | -105.1385
19.0498 | -105.141
19.0536 | -105.1428
19.0573 | -105.1551
19.0611 | -105.1609
19.0649 | -105.1681
19.0686 | -105.176
19.0724 | -105.1862
19.0762 | -105.1967
19.08 | -105.2068
19.0837 | -105.2181
19.0875 | -105.2331
19.0913 | -105.2483
19.0951 | -105.2616
19.0988 | -105.2795
19.1026 | -105.2988
19.1064 | -105.3175
19.1102 | -105.3338
19.1139 | -105.3506
19.1177 | -105.3692
19.1215 | -105.3857
19.1253 | -105.3987
19.129 | -105.4107
19.1328 | -105.4235
19.1366 | -105.4369
19.1404 | -105.4481
19.1441 | -105.4591
19.1479 | -105.4699
19.1517 | -105.4795
19.1555 | -105.4876
19.1592 | -105.4926
19.163 | -105.4924
19.1668 | -105.4907
19.1706 | -105.4889
19.1743 | -105.4861
19.1781 | -105.4838
19.1819 | -105.4831
19.1857 | -105.4784
19.1894 | -105.4714
19.1932 | -105.4614
19.197 | -105.4522
19.2008 | -105.4415
19.2045 | -105.4295
19.2083 | -105.418
19.2121 | -105.4082
19.2159 | -105.3969
19.2196 | -105.386
19.2234 | -105.374
19.2272 | -105.3605
19.231 | -105.3468
19.2347 | -105.3344
19.2385 | -105.3222
19.2423 | -105.3074
19.246 | -105.2915
19.2498 | -105.2772
19.2536 | -105.2651
19.2574 | -105.2531
19.2611 | -105.2395
19.2649 | -105.228
19.2687 | -105.2195
19.2725 | -105.2125
19.2762 | -105.2035
19.28 | -105.1953
19.2838 | -105.1875
19.2876 | -105.1802
19.2913 | -105.1737
19.2951 | -105.1686
19.2989 | -105.164
19.3027 | -105.1589
19.3064 | -105.1574
19.3102 | -105.1541
19.314 | -105.1523
19.3178 | -105.1516
19.3215 | -105.1496
19.3253 | -105.1494
19.3291 | -105.1506
19.3329 | -105.1522
19.3366 | -105.153
19.3404 | -105.1531
19.3442 | -105.1509
19.348 | -105.1503
19.3517 | -105.151
19.3555 | -105.1534
19.3593 | -105.155
19.3631 | -105.1561
19.3668 | -105.159
19.3706 | -105.1639
19.3744 | -105.1672
19.3782 | -105.1699
19.3819 | -105.1697
19.3857 | -105.1697
19.3895 | -105.1703
19.3933 | -105.1695
19.397 | -105.1702
19.4008 | -105.1723
19.4046 | -105.1728
19.4084 | -105.1732
19.4121 | -105.1717
19.4159 | -105.1718
19.4197 | -105.1692
19.4234 | -105.1659
19.4272 | -105.1614
19.431 | -105.1586
19.4348 | -105.1572
19.4385 | -105.1562
19.4423 | -105.152
19.4461 | -105.1478
19.4499 | -105.1432
19.4536 | -105.1373
19.4574 | -105.1324
19.4612 | -105.1288
19.465 | -105.1232
19.4687 | -105.1196
19.4725 | -105.1172
19.4763 | -105.1118
19.4801 | -105.1061
19.4838 | -105.1026
19.4876 | -105.0973
19.4914 | -105.0931
19.4952 | -105.087
19.4989 | -105.082
19.5027 | -105.0792
19.5065 | -105.0762
19.5103 | -105.0733
19.514 | -105.0736
19.5178 | -105.073
19.5216 | -105.0732
19.5254 | -105.0733
19.5291 | -105.0719
19.5329 | -105.0712
19.5367 | -105.0716
19.5405 | -105.0709
19.5442 | -105.0665
19.548 | -105.0636
19.5518 | -105.0629
19.5556 | -105.0629
19.5593 | -105.0642
19.5631 | -105.0642
19.5669 | -105.0636
19.5707 | -105.0623
19.5744 | -105.0613
19.5782 | -105.0619
19.582 | -105.0594
19.5858 | -105.0576
19.5895 | -105.0546
19.5933 | -105.0492
19.5971 | -105.0469
19.6008 | -105.0402
19.6046 | -105.0358
19.6084 | -105.0315
19.6122 | -105.027
19.6159 | -105.0236
19.6197 | -105.0205
19.6235 | -105.0192
19.6273 | -105.0165
19.631 | -105.0133
19.6348 | -105.0096
19.6386 | -105.0079
19.6424 | -105.0062
19.6461 | -105.0065
19.6499 | -105.0071
19.6537 | -105.0066
19.6575 | -105.0077
19.6612 | -105.0125
19.665 | -105.0182
19.6688 | -105.0235
19.6726 | -105.0298
19.6763 | -105.0365
19.6801 | -105.0454
19.6839 | -105.0548
19.6877 | -105.0641
19.6914 | -105.0745
19.6952 | -105.0844
19.699 | -105.0966
19.7028 | -105.1071
19.7065 | -105.1205
19.7103 | -105.1336
19.7141 | -105.1462
19.7179 | -105.1605
19.7216 | -105.1744
19.7254 | -105.1884
19.7292 | -105.2027
19.733 | -105.2176
19.7367 | -105.2295
19.7405 | -105.2399
19.7443 | -105.2517
19.7481 | -105.2636
19.7518 | -105.2729
19.7556 | -105.2827
19.7594 | -105.2935
19.7632 | -105.3013
19.7669 | -105.3077
19.7707 | -105.313
19.7745 | -105.3193
19.7782 | -105.324
19.782 | -105.3274
19.7858 | -105.3292
19.7896 | -105.3295
19.7933 | -105.3284
19.7971 | -105.3275
19.8009 | -105.3253
19.8047 | -105.3221
19.8084 | -105.3183
19.8122 | -105.314
19.816 | -105.3086
19.8198 | -105.3044
19.8235 | -105.3011
19.8273 | -105.2961
19.8311 | -105.2886
19.8349 | -105.279
19.8386 | -105.2689
19.8424 | -105.2598
19.8462 | -105.2512
19.85 | -105.2412
19.8537 | -105.2298
19.8575 | -105.22
19.8613 | -105.2103
19.8651 | -105.2004
19.8688 | -105.1924
19.8726 | -105.1848
19.8764 | -105.1757
19.8802 | -105.1663
19.8839 | -105.1584
19.8877 | -105.1518
19.8915 | -105.144
19.8953 | -105.135
19.899 | -105.1277
19.9028 | -105.1211
19.9066 | -105.1138
19.9104 | -105.1061
19.9141 | -105.0988
19.9179 | -105.0942
19.9217 | -105.0882
19.9255 | -105.0832
19.9292 | -105.0772
19.933 | -105.0707
19.9368 | -105.0686
19.9406 | -105.0656
19.9443 | -105.0599
19.9481 | -105.0563
19.9519 | -105.0543
19.9556 | -105.0526
19.9594 | -105.0513
19.9632 | -105.0486
19.967 | -105.0461
19.9707 | -105.0419
19.9745 | -105.04
19.9783 | -105.0387
19.9821 | -105.039
19.9858 | -105.0393
19.9896 | -105.0369
19.9934 | -105.0338
19.9972 | -105.0303
20.0009 | -105.026
20.0047 | -105.0255
20.0085 | -105.0213
20.0123 | -105.0158
20.016 | -105.0104
20.0198 | -105.0059
20.0236 | -105.0023
20.0274 | -104.9982
20.0311 | -104.9925
20.0349 | -104.988
20.0387 | -104.9838
20.0425 | -104.9788
20.0462 | -104.9737
20.05 | -104.9701
20.0538 | -104.9647
20.0576 | -104.9578
20.0613 | -104.9522
20.0651 | -104.9452
20.0689 | -104.9385
20.0727 | -104.9338
20.0764 | -104.9288
20.0802 | -104.9239
20.084 | -104.9196
20.0878 | -104.9167
20.0915 | -104.9112
20.0953 | -104.9068
20.0991 | -104.9045
20.1029 | -104.9021
20.1066 | -104.8996
20.1104 | -104.8968
20.1142 | -104.8959
20.118 | -104.8951
20.1217 | -104.8956
20.1255 | -104.8977
20.1293 | -104.8986
20.1331 | -104.8993
20.1368 | -104.9026
20.1406 | -104.9024
20.1444 | -104.8982
20.1481 | -104.898
20.1519 | -104.8979
20.1557 | -104.8966
20.1595 | -104.895
20.1632 | -104.894
20.167 | -104.8935
20.1708 | -104.8922
20.1746 | -104.8885
20.1783 | -104.8863
20.1821 | -104.8857
20.1859 | -104.8831
20.1897 | -104.8807
20.1934 | -104.8759
20.1972 | -104.8724
20.201 | -104.8695
20.2048 | -104.8618
20.2085 | -104.8552
20.2123 | -104.8486
20.2161 | -104.8425
20.2199 | -104.8372
20.2236 | -104.8318
20.2274 | -104.8266
20.2312 | -104.822
20.235 | -104.8154
20.2387 | -104.8081
20.2425 | -104.8013
20.2463 | -104.7946
20.2501 | -104.7906
20.2538 | -104.7859
20.2576 | -104.7838
20.2614 | -104.781
20.2652 | -104.7779
20.2689 | -104.7756
20.2727 | -104.7762
20.2765 | -104.7788
20.2803 | -104.7826
20.284 | -104.7849
20.2878 | -104.7873
20.2916 | -104.7902
20.2954 | -104.7959
20.2991 | -104.8018
20.3029 | -104.8105
20.3067 | -104.8193
20.3105 | -104.8277
20.3142 | -104.8351
20.318 | -104.8447
20.3218 | -104.8549
20.3255 | -104.8665
20.3293 | -104.8776
20.3331 | -104.8894
20.3369 | -104.8991
20.3406 | -104.9104
20.3444 | -104.9237
20.3482 | -104.937
20.352 | -104.9496
20.3557 | -104.9611
20.3595 | -104.9701
20.3633 | -104.979
20.3671 | -104.9885
20.3708 | -104.9969
20.3746 | -105.0033
20.3784 | -105.0108
20.3822 | -105.0173
20.3859 | -105.0222
20.3897 | -105.0264
20.3935 | -105.0299
20.3973 | -105.0338
20.401 | -105.037
20.4048 | -105.0403
20.4086 | -105.0406
20.4124 | -105.0387
20.4161 | -105.0388
20.4199 | -105.0391
20.4237 | -105.0344
20.4275 | -105.0337
20.4312 | -105.0347
20.435 | -105.0344
20.4388 | -105.0301
20.4426 | -105.025
20.4463 | -105.0225
20.4501 | -105.0204
20.4539 | -105.0156
20.4577 | -105.0121
20.4614 | -105.0087
20.4652 | -105.0074
20.469 | -105.0026
20.4728 | -104.9972
20.4765 | -104.9946
20.4803 | -104.9916
20.4841 | -104.9924
20.4879 | -104.9916
20.4916 | -104.9881
20.4954 | -104.9868
20.4992 | -104.9853
20.5029 | -104.9837
20.5067 | -104.9806
20.5105 | -104.9772
20.5143 | -104.9718
20.518 | -104.967
20.5218 | -104.9648
20.5256 | -104.9633
20.5294 | -104.962
20.5331 | -104.9604
20.5369 | -104.9595
20.5407 | -104.9567
20.5445 | -104.9549
20.5482 | -104.9523
20.552 | -104.9494
20.5558 | -104.9475
20.5596 | -104.9445
20.5633 | -104.9415
20.5671 | -104.9379
20.5709 | -104.9346
20.5747 | -104.933
20.5784 | -104.9312
20.5822 | -104.9268
20.586 | -104.9229
20.5898 | -104.9213
20.5935 | -104.9178
20.5973 | -104.9153
20.6011 | -104.9122
20.6049 | -104.9111
20.6086 | -104.9091
20.6124 | -104.9065
20.6162 | -104.9022
20.62 | -104.901
20.6237 | -104.8976
20.6275 | -104.8951
20.6313 | -104.893
20.6351 | -104.891
20.6388 | -104.8868
20.6426 | -104.8831
20.6464 | -104.8833
20.6502 | -104.8837
20.6539 | -104.8821
20.6577 | -104.8787
20.6615 | -104.8753
20.6653 | -104.8725
20.669 | -104.8714
20.6728 | -104.8705
20.6766 | -104.869
20.6803 | -104.8659
20.6841 | -104.8627
20.6879 | -104.8592
20.6917 | -104.8539
20.6954 | -104.8479
20.6992 | -104.8431
20.703 | -104.8388
20.7068 | -104.8352
20.7105 | -104.8326
20.7143 | -104.8298
20.7181 | -104.8273
20.7219 | -104.8235
20.7256 | -104.8183
20.7294 | -104.8153
20.7332 | -104.8123
20.737 | -104.8083
20.7407 | -104.8026
20.7445 | -104.7975
20.7483 | -104.7932
20.7521 | -104.7888
20.7558 | -104.7857
20.7596 | -104.7818
20.7634 | -104.779
20.7672 | -104.7767
20.7709 | -104.7727
20.7747 | -104.7697
20.7785 | -104.7649
20.7823 | -104.7611
20.786 | -104.7604
20.7898 | -104.7599
20.7936 | -104.7591
20.7974 | -104.7569
20.8011 | -104.7566
20.8049 | -104.7574
20.8087 | -104.7587
20.8125 | -104.7589
20.8162 | -104.7566
20.82 | -104.7568
20.8238 | -104.7595
20.8276 | -104.7628
20.8313 | -104.7657
20.8351 | -104.7679
20.8389 | -104.7699
20.8427 | -104.7734
20.8464 | -104.7786
20.8502 | -104.7834
20.854 | -104.7892
20.8577 | -104.7953
20.8615 | -104.8034
20.8653 | -104.8112
20.8691 | -104.8187
20.8728 | -104.8274
20.8766 | -104.8367
20.8804 | -104.8463
20.8842 | -104.853
20.8879 | -104.8604
20.8917 | -104.8682
20.8955 | -104.8734
20.8993 | -104.8788
20.903 | -104.8843
20.9068 | -104.8903
20.9106 | -104.8979
20.9144 | -104.9026
20.9181 | -104.9093
20.9219 | -104.9165
20.9257 | -104.9219
20.9295 | -104.9251
20.9332 | -104.9296
20.937 | -104.9325
20.9408 | -104.9308
20.9446 | -104.9319
20.9483 | -104.9345
20.9521 | -104.9369
20.9559 | -104.939
20.9597 | -104.9418
20.9634 | -104.9407
20.9672 | -104.9374
20.971 | -104.9352
20.9748 | -104.9344
20.9785 | -104.9341
20.9823 | -104.9318
20.9861 | -104.9292
20.9899 | -104.9253
20.9936 | -104.9234
20.9974 | -104.9194
21.0012 | -104.9146
21.005 | -104.9101
21.0087 | -104.9029
21.0125 | -104.8942
21.0163 | -104.8886
21.0201 | -104.8842
21.0238 | -104.8788
21.0276 | -104.8729
21.0314 | -104.8698
21.0352 | -104.8673
21.0389 | -104.8641
21.0427 | -104.8607
21.0465 | -104.8593
21.0502 | -104.8571
21.054 | -104.8559
21.0578 | -104.8552
21.0616 | -104.8549
21.0653 | -104.8536
21.0691 | -104.8542
21.0729 | -104.8578
21.0767 | -104.8607
21.0804 | -104.8619
21.0842 | -104.8638
21.088 | -104.8654
21.0918 | -104.8666
21.0955 | -104.8681
21.0993 | -104.8683
21.1031 | -104.87
21.1069 | -104.8727
21.1106 | -104.876
21.1144 | -104.8788
21.1182 | -104.8792
21.122 | -104.8772
21.1257 | -104.8772
21.1295 | -104.8779
21.1333 | -104.8785
21.1371 | -104.8804
21.1408 | -104.8805
21.1446 | -104.8773
21.1484 | -104.8754
21.1522 | -104.8736
21.1559 | -104.8711
21.1597 | -104.8652
21.1635 | -104.8605
21.1673 | -104.8573
21.171 | -104.8534
21.1748 | -104.8491
21.1786 | -104.8438
21.1824 | -104.8401
21.1861 | -104.8369
21.1899 | -104.8327
21.1937 | -104.8329
21.1975 | -104.8313
21.2012 | -104.83
21.205 | -104.8261
21.2088 | -104.8219
21.2126 | -104.8213
21.2163 | -104.822
21.2201 | -104.8186
21.2239 | -104.8179
21.2276 | -104.8198
21.2314 | -104.8214
21.2352 | -104.8258
21.239 | -104.8285
21.2427 | -104.8316
21.2465 | -104.8327
21.2503 | -104.8342
21.2541 | -104.8362
21.2578 | -104.8379
21.2616 | -104.8393
21.2654 | -104.8427
21.2692 | -104.8473
21.2729 | -104.8523
21.2767 | -104.8564
21.2805 | -104.8606
21.2843 | -104.8668
21.288 | -104.8715
21.2918 | -104.8752
21.2956 | -104.8779
21.2994 | -104.8818
21.3031 | -104.8856
21.3069 | -104.8886
21.3107 | -104.8928
21.3145 | -104.8994
21.3182 | -104.9041
21.322 | -104.9075
21.3258 | -104.9114
21.3296 | -104.9158
21.3333 | -104.919
21.3371 | -104.9229
21.3409 | -104.9254
21.3447 | -104.9289
21.3484 | -104.9297
21.3522 | -104.9301
21.356 | -104.9294
21.3598 | -104.9307
21.3635 | -104.9332
21.3673 | -104.9361
21.3711 | -104.9366
21.3749 | -104.9351
21.3786 | -104.935
21.3824 | -104.9361
21.3862 | -104.9352
21.3899 | -104.9337
21.3937 | -104.9326
21.3975 | -104.9319
21.4013 | -104.9318
21.405 | -104.9296
21.4088 | -104.9252
21.4126 | -104.92
21.4164 | -104.9164
21.4201 | -104.9111
21.4239 | -104.9044
21.4277 | -104.897
21.4315 | -104.8902
21.4352 | -104.8826
21.439 | -104.8739
21.4428 | -104.8652
21.4466 | -104.8598
21.4503 | -104.8527
21.4541 | -104.8427
21.4579 | -104.8335
21.4617 | -104.8237
21.4654 | -104.8161
21.4692 | -104.8087
21.473 | -104.8016
21.4768 | -104.794
21.4805 | -104.7849
21.4843 | -104.7764
21.4881 | -104.7678
21.4919 | -104.7583
21.4956 | -104.7489
21.4994 | -104.7396
21.5032 | -104.7307
21.507 | -104.7252
21.5107 | -104.7196
21.5145 | -104.7112
21.5183 | -104.7043
21.5221 | -104.6983
21.5258 | -104.6916
21.5296 | -104.6831
21.5334 | -104.6748
21.5372 | -104.6673
21.5409 | -104.6623
21.5447 | -104.6564
21.5485 | -104.6524
21.5523 | -104.6469
21.556 | -104.6409
21.5598 | -104.6346
21.5636 | -104.6288
21.5674 | -104.6232
21.5711 | -104.6176
21.5749 | -104.6113
21.5787 | -104.6042
21.5824 | -104.5988
21.5862 | -104.5962
21.59 | -104.5915
21.5938 | -104.5873
21.5975 | -104.5852
21.6013 | -104.5815
21.6051 | -104.5768
21.6089 | -104.5747
21.6126 | -104.5722
21.6164 | -104.569
21.6202 | -104.5635
21.624 | -104.558
21.6277 | -104.5549
21.6315 | -104.5535
21.6353 | -104.5506
21.6391 | -104.5494
21.6428 | -104.5472
21.6466 | -104.5462
21.6504 | -104.5461
21.6542 | -104.548
21.6579 | -104.548
21.6617 | -104.5461
21.6655 | -104.5447
21.6693 | -104.5451
21.673 | -104.5457
21.6768 | -104.5461
21.6806 | -104.5464
21.6844 | -104.5458
21.6881 | -104.546
21.6919 | -104.5476
21.6957 | -104.5479
21.6995 | -104.5472
21.7032 | -104.5474
21.707 | -104.5472
21.7108 | -104.5456
21.7146 | -104.5468
21.7183 | -104.5513
21.7221 | -104.5548
21.7259 | -104.5585
21.7297 | -104.5628
21.7334 | -104.5669
21.7372 | -104.57
21.741 | -104.5737
21.7448 | -104.5763
21.7485 | -104.5816
21.7523 | -104.5846
21.7561 | -104.5878
21.7598 | -104.5924
21.7636 | -104.5958
21.7674 | -104.6023
21.7712 | -104.6056
21.7749 | -104.6092
21.7787 | -104.6117
21.7825 | -104.6144
21.7863 | -104.6176
21.79 | -104.6232
21.7938 | -104.6284
21.7976 | -104.6313
21.8014 | -104.6348
21.8051 | -104.6373
21.8089 | -104.6405
21.8127 | -104.6434
21.8165 | -104.6447
21.8202 | -104.6465
21.824 | -104.6463
21.8278 | -104.6455
21.8316 | -104.6475
21.8353 | -104.6497
21.8391 | -104.6525
21.8429 | -104.6542
21.8467 | -104.654
21.8504 | -104.6519
21.8542 | -104.6524
21.858 | -104.6516
21.8618 | -104.6523
21.8655 | -104.6517
21.8693 | -104.6505
21.8731 | -104.6488
21.8769 | -104.6448
21.8806 | -104.643
21.8844 | -104.642
21.8882 | -104.6393
21.892 | -104.6357
21.8957 | -104.6313
21.8995 | -104.6299
21.9033 | -104.6289
21.9071 | -104.6289
21.9108 | -104.6301
21.9146 | -104.6305
21.9184 | -104.6301
21.9222 | -104.6317
21.9259 | -104.6334
21.9297 | -104.6362
21.9335 | -104.6395
21.9372 | -104.641
21.941 | -104.6422
21.9448 | -104.6451
21.9486 | -104.647
21.9523 | -104.6509
21.9561 | -104.6556
21.9599 | -104.6582
21.9637 | -104.6601
21.9674 | -104.6665
21.9712 | -104.6734
21.975 | -104.6765
21.9788 | -104.6785
21.9825 | -104.6805
21.9863 | -104.683
21.9901 | -104.6875
21.9939 | -104.6937
21.9976 | -104.6984
22.0014 | -104.7
22.0052 | -104.6993
22.009 | -104.6977
22.0127 | -104.6973
22.0165 | -104.6954
22.0203 | -104.6928
22.0241 | -104.6871
22.0278 | -104.6793
22.0316 | -104.6723
22.0354 | -104.665
22.0392 | -104.6559
22.0429 | -104.6466
22.0467 | -104.6386
22.0505 | -104.63
22.0543 | -104.621
22.058 | -104.6088
22.0618 | -104.5957
22.0656 | -104.5827
22.0694 | -104.5713
22.0731 | -104.561
22.0769 | -104.5492
22.0807 | -104.5375
22.0845 | -104.5235
22.0882 | -104.5115
22.092 | -104.4995
22.0958 | -104.4866
22.0996 | -104.474
22.1033 | -104.4611
22.1071 | -104.4475
22.1109 | -104.4327
22.1147 | -104.4216
22.1184 | -104.4129
22.1222 | -104.405
22.126 | -104.3959
22.1297 | -104.3874
22.1335 | -104.3807
22.1373 | -104.3753
22.1411 | -104.3697
22.1448 | -104.3665
22.1486 | -104.3631
22.1524 | -104.3597
22.1562 | -104.3571
22.1599 | -104.3555
22.1637 | -104.355
22.1675 | -104.3549
22.1713 | -104.3568
22.175 | -104.3598
22.1788 | -104.3653
22.1826 | -104.3699
22.1864 | -104.3748
22.1901 | -104.3813
22.1939 | -104.3867
22.1977 | -104.393
22.2015 | -104.4007
22.2052 | -104.4086
22.209 | -104.4161
22.2128 | -104.4244
22.2166 | -104.4338
22.2203 | -104.4433
22.2241 | -104.4537
22.2279 | -104.4661
22.2317 | -104.4764
22.2354 | -104.4852
22.2392 | -104.4962
22.243 | -104.5084
22.2468 | -104.5214
22.2505 | -104.5339
22.2543 | -104.5462
22.2581 | -104.558
22.2619 | -104.5701
22.2656 | -104.5818
22.2694 | -104.5927
22.2732 | -104.6029
22.277 | -104.6118
22.2807 | -104.6196
22.2845 | -104.6282
22.2883 | -104.6358
22.292 | -104.6445
22.2958 | -104.6533
22.2996 | -104.6586
22.3034 | -104.6599
22.3071 | -104.6637
22.3109 | -104.6664
22.3147 | -104.6707
22.3185 | -104.6769
22.3222 | -104.6822
22.326 | -104.6819
22.3298 | -104.6847
22.3336 | -104.6868
22.3373 | -104.6894
22.3411 | -104.6921
22.3449 | -104.6955
22.3487 | -104.6971
22.3524 | -104.696
22.3562 | -104.6945
22.36 | -104.6908
22.3638 | -104.6882
22.3675 | -104.6852
22.3713 | -104.682
22.3751 | -104.677
22.3789 | -104.6746
22.3826 | -104.6712
22.3864 | -104.6645
22.3902 | -104.6585
22.394 | -104.6505
22.3977 | -104.6451
22.4015 | -104.6409
22.4053 | -104.6352
22.4091 | -104.6298
22.4128 | -104.6248
22.4166 | -104.6204
22.4204 | -104.6138
22.4242 | -104.6078
22.4279 | -104.6059
22.4317 | -104.6002
22.4355 | -104.5917
22.4393 | -104.5828
22.443 | -104.5753
22.4468 | -104.5692
22.4506 | -104.5612
22.4544 | -104.5546
22.4581 | -104.5492
22.4619 | -104.5391
22.4657 | -104.5292
22.4694 | -104.5227
22.4732 | -104.5179
22.477 | -104.5132
22.4808 | -104.5073
22.4845 | -104.4996
22.4883 | -104.4928
22.4921 | -104.4861
22.4959 | -104.4787
22.4996 | -104.4721
22.5034 | -104.4693
22.5072 | -104.4661
22.511 | -104.4618
22.5147 | -104.4576
22.5185 | -104.4528
22.5223 | -104.4492
22.5261 | -104.4452
22.5298 | -104.4402
22.5336 | -104.4361
22.5374 | -104.4309
22.5412 | -104.4246
22.5449 | -104.4189
22.5487 | -104.4138
22.5525 | -104.41
22.5563 | -104.4055
22.56 | -104.4018
22.5638 | -104.3961
22.5676 | -104.3901
22.5714 | -104.3843
22.5751 | -104.3781
22.5789 | -104.3718
22.5827 | -104.3655
22.5865 | -104.3603
22.5902 | -104.3568
22.594 | -104.354
22.5978 | -104.349
22.6016 | -104.3441
22.6053 | -104.3417
22.6091 | -104.3376
22.6129 | -104.332
22.6167 | -104.3301
22.6204 | -104.3285
22.6242 | -104.3226
22.628 | -104.3177
22.6318 | -104.3119
22.6355 | -104.3068
22.6393 | -104.3031
22.6431 | -104.2981
22.6469 | -104.2923
22.6506 | -104.2882
22.6544 | -104.2833
22.6582 | -104.2768
22.6619 | -104.2712
22.6657 | -104.2647
22.6695 | -104.2569
22.6733 | -104.2503
22.677 | -104.2458
22.6808 | -104.2426
22.6846 | -104.2364
22.6884 | -104.2304
22.6921 | -104.2241
22.6959 | -104.2194
22.6997 | -104.2138
22.7035 | -104.2084
22.7072 | -104.2058
22.711 | -104.1998
22.7148 | -104.1955
22.7186 | -104.1902
22.7223 | -104.1845
22.7261 | -104.1779
22.7299 | -104.1711
22.7337 | -104.1644
22.7374 | -104.1585
22.7412 | -104.1545
22.745 | -104.15
22.7488 | -104.144
22.7525 | -104.1404
22.7563 | -104.1371
22.7601 | -104.1356
22.7639 | -104.1338
22.7676 | -104.1315
22.7714 | -104.128
22.7752 | -104.1237
22.779 | -104.1182
22.7827 | -104.1164
22.7865 | -104.1147
22.7903 | -104.1112
22.7941 | -104.1073
22.7978 | -104.1053
22.8016 | -104.1019
22.8054 | -104.1
22.8092 | -104.0992
22.8129 | -104.0996
22.8167 | -104.1
22.8205 | -104.1022
22.8243 | -104.1016
22.828 | -104.0998
22.8318 | -104.0996
22.8356 | -104.1003
22.8393 | -104.1017
22.8431 | -104.1015
22.8469 | -104.102
22.8507 | -104.1022
22.8544 | -104.1052
22.8582 | -104.1085
22.862 | -104.1115
22.8658 | -104.1142
22.8695 | -104.1162
22.8733 | -104.1169
22.8771 | -104.1173
22.8809 | -104.1188
22.8846 | -104.119
22.8884 | -104.1209
22.8922 | -104.1243
22.896 | -104.1265
22.8997 | -104.1279
22.9035 | -104.1289
22.9073 | -104.1288
22.9111 | -104.128
22.9148 | -104.1268
22.9186 | -104.1266
22.9224 | -104.1245
22.9262 | -104.12
22.9299 | -104.116
22.9337 | -104.1129
22.9375 | -104.1087
22.9413 | -104.1038
22.945 | -104.1001
22.9488 | -104.0956
22.9526 | -104.0907
22.9564 | -104.0874
22.9601 | -104.0848
22.9639 | -104.0796
22.9677 | -104.0756
22.9715 | -104.0705
22.9752 | -104.0653
22.979 | -104.0622
22.9828 | -104.0612
22.9866 | -104.0569
22.9903 | -104.0543
22.9941 | -104.0549
22.9979 | -104.0547
23.0017 | -104.0545
23.0054 | -104.0548
23.0092 | -104.0557
23.013 | -104.0571
23.0167 | -104.058
23.0205 | -104.0617
23.0243 | -104.0649
23.0281 | -104.069
23.0318 | -104.0727
23.0356 | -104.0777
23.0394 | -104.0809
23.0432 | -104.0847
23.0469 | -104.091
23.0507 | -104.0971
23.0545 | -104.1047
23.0583 | -104.1124
23.062 | -104.1171
23.0658 | -104.1214
23.0696 | -104.1267
23.0734 | -104.133
23.0771 | -104.1383
23.0809 | -104.143
23.0847 | -104.1465
23.0885 | -104.1508
23.0922 | -104.158
23.096 | -104.1675
23.0998 | -104.1756
23.1036 | -104.1828
23.1073 | -104.1876
23.1111 | -104.192
23.1149 | -104.1971
23.1187 | -104.2014
23.1224 | -104.2032
23.1262 | -104.2066
23.13 | -104.2088
23.1338 | -104.2085
23.1375 | -104.2085
23.1413 | -104.2097
23.1451 | -104.2106
23.1489 | -104.211
23.1526 | -104.2131
23.1564 | -104.2137
23.1602 | -104.2131
23.164 | -104.2127
23.1677 | -104.2126
23.1715 | -104.2123
23.1753 | -104.2123
23.1791 | -104.2108
23.1828 | -104.2102
23.1866 | -104.2116
23.1904 | -104.2116
23.1942 | -104.2093
23.1979 | -104.2067
23.2017 | -104.2069
23.2055 | -104.2067
23.2092 | -104.2049
23.213 | -104.2019
23.2168 | -104.2012
23.2206 | -104.2001
23.2243 | -104.1986
23.2281 | -104.1979
23.2319 | -104.1984
23.2357 | -104.2001
23.2394 | -104.2015
23.2432 | -104.2023
23.247 | -104.2033
23.2508 | -104.2076
23.2545 | -104.2147
23.2583 | -104.2202
23.2621 | -104.2244
23.2659 | -104.2303
23.2696 | -104.2353
23.2734 | -104.2404
23.2772 | -104.2488
23.281 | -104.2572
23.2847 | -104.2638
23.2885 | -104.2707
23.2923 | -104.2791
23.2961 | -104.2874
23.2998 | -104.2972
23.3036 | -104.3075
23.3074 | -104.3185
23.3112 | -104.3278
23.3149 | -104.3398
23.3187 | -104.3512
23.3225 | -104.3614
23.3263 | -104.3736
23.33 | -104.3839
23.3338 | -104.3938
23.3376 | -104.4044
23.3414 | -104.4171
23.3451 | -104.4309
23.3489 | -104.4408
23.3527 | -104.4492
23.3565 | -104.4582
23.3602 | -104.4663
23.364 | -104.4772
23.3678 | -104.4897
23.3715 | -104.502
23.3753 | -104.5133
23.3791 | -104.527
23.3829 | -104.5372
23.3866 | -104.5462
23.3904 | -104.5582
23.3942 | -104.5699
23.398 | -104.5791
23.4017 | -104.587
23.4055 | -104.5952
23.4093 | -104.6043
23.4131 | -104.6084
23.4168 | -104.6166
23.4206 | -104.6244
23.4244 | -104.6273
23.4282 | -104.6287
23.4319 | -104.631
23.4357 | -104.6349
23.4395 | -104.6388
23.4433 | -104.6415
23.447 | -104.6419
23.4508 | -104.6437
23.4546 | -104.6469
23.4584 | -104.6492
23.4621 | -104.6516
23.4659 | -104.653
23.4697 | -104.6532
23.4735 | -104.6508
23.4772 | -104.6477
23.481 | -104.645
23.4848 | -104.643
23.4886 | -104.6422
23.4923 | -104.6405
23.4961 | -104.6399
23.4999 | -104.6396
23.5037 | -104.6393
23.5074 | -104.638
23.5112 | -104.6349
23.515 | -104.6311
23.5188 | -104.6274
23.5225 | -104.626
23.5263 | -104.6264
23.5301 | -104.626
23.5339 | -104.6251
23.5376 | -104.624
23.5414 | -104.623
23.5452 | -104.6241
23.549 | -104.6227
23.5527 | -104.6207
23.5565 | -104.6206
23.5603 | -104.6209
23.564 | -104.6195
23.5678 | -104.6178
23.5716 | -104.6149
23.5754 | -104.6112
23.5791 | -104.6062
23.5829 | -104.6027
23.5867 | -104.6008
23.5905 | -104.5998
23.5942 | -104.5974
23.598 | -104.5956
23.6018 | -104.5943
23.6056 | -104.5943
23.6093 | -104.5929
23.6131 | -104.5913
23.6169 | -104.5912
23.6207 | -104.5909
23.6244 | -104.5922
23.6282 | -104.5938
23.632 | -104.5954
23.6358 | -104.5981
23.6395 | -104.601
23.6433 | -104.6047
23.6471 | -104.609
23.6509 | -104.613
23.6546 | -104.6192
23.6584 | -104.6255
23.6622 | -104.6309
23.666 | -104.6371
23.6697 | -104.6443
23.6735 | -104.65
23.6773 | -104.6565
23.6811 | -104.6649
23.6848 | -104.6745
23.6886 | -104.6824
23.6924 | -104.6905
23.6962 | -104.6979
23.6999 | -104.7042
23.7037 | -104.7107
23.7075 | -104.7164
23.7113 | -104.723
23.715 | -104.7317
23.7188 | -104.7406
23.7226 | -104.7477
23.7264 | -104.7548
23.7301 | -104.7631
23.7339 | -104.7691
23.7377 | -104.7735
23.7414 | -104.78
23.7452 | -104.7863
23.749 | -104.7918
23.7528 | -104.7973
23.7565 | -104.8013
23.7603 | -104.8082
23.7641 | -104.8144
23.7679 | -104.8203
23.7716 | -104.8216
23.7754 | -104.8245
23.7792 | -104.8281
23.783 | -104.8324
23.7867 | -104.8329
23.7905 | -104.8323
23.7943 | -104.8321
23.7981 | -104.83
23.8018 | -104.8263
23.8056 | -104.8224
23.8094 | -104.8202
23.8132 | -104.8161
23.8169 | -104.8134
23.8207 | -104.8107
23.8245 | -104.8055
23.8283 | -104.7984
23.832 | -104.7913
23.8358 | -104.7834
23.8396 | -104.7746
23.8434 | -104.767
23.8471 | -104.7601
23.8509 | -104.7537
23.8547 | -104.7453
23.8585 | -104.736
23.8622 | -104.7271
23.866 | -104.7182
23.8698 | -104.7082
23.8736 | -104.6987
23.8773 | -104.6908
23.8811 | -104.6822
23.8849 | -104.6723
23.8887 | -104.6624
23.8924 | -104.6535
23.8962 | -104.6447
23.9 | -104.6359
23.9038 | -104.6289
23.9075 | -104.6235
23.9113 | -104.6196
23.9151 | -104.616
23.9188 | -104.6113
23.9226 | -104.6056
23.9264 | -104.6014
23.9302 | -104.5989
23.9339 | -104.5971
23.9377 | -104.5962
23.9415 | -104.5947
23.9453 | -104.5944
23.949 | -104.5953
23.9528 | -104.5956
23.9566 | -104.6038
23.9604 | -104.6054
23.9641 | -104.6064
23.9679 | -104.6091
23.9717 | -104.6124
23.9755 | -104.6162
23.9792 | -104.6208
23.983 | -104.6266
23.9868 | -104.6322
23.9906 | -104.637
23.9943 | -104.6411
23.9981 | -104.6479
24.0019 | -104.6571
24.0057 | -104.6656
24.0094 | -104.6732
24.0132 | -104.6822
24.017 | -104.6911
24.0208 | -104.7001
24.0245 | -104.7105
24.0283 | -104.7221
24.0321 | -104.735
24.0359 | -104.7463
24.0396 | -104.7574
24.0434 | -104.7669
24.0472 | -104.7766
24.051 | -104.7859
24.0547 | -104.7955
24.0585 | -104.805
24.0623 | -104.8137
24.0661 | -104.8227
24.0698 | -104.8316
24.0736 | -104.8414
24.0774 | -104.8498
24.0812 | -104.8563
24.0849 | -104.8632
24.0887 | -104.8715
24.0925 | -104.8786
24.0962 | -104.8857
24.1 | -104.8898
24.1038 | -104.8954
24.1076 | -104.9032
24.1113 | -104.9122
24.1151 | -104.9184
24.1189 | -104.9216
24.1227 | -104.9243
24.1264 | -104.9279
24.1302 | -104.932
24.134 | -104.9332
24.1378 | -104.9355
24.1415 | -104.9377
24.1453 | -104.9387
24.1491 | -104.9391
24.1529 | -104.9404
24.1566 | -104.939
24.1604 | -104.9357
24.1642 | -104.9331
24.168 | -104.9316
24.1717 | -104.9297
24.1755 | -104.9277
24.1793 | -104.9245
24.1831 | -104.9193
24.1868 | -104.9146
24.1906 | -104.9091
24.1944 | -104.9041
24.1982 | -104.898
24.2019 | -104.8923
24.2057 | -104.8884
24.2095 | -104.8838
24.2133 | -104.8773
24.217 | -104.8684
24.2208 | -104.8596
24.2246 | -104.8521
24.2284 | -104.846
24.2321 | -104.8393
24.2359 | -104.8314
24.2397 | -104.8235
24.2435 | -104.8171
24.2472 | -104.8105
24.251 | -104.803
24.2548 | -104.7946
24.2586 | -104.7845
24.2623 | -104.775
24.2661 | -104.7641
24.2699 | -104.7544
24.2736 | -104.744
24.2774 | -104.7324
24.2812 | -104.7197
24.285 | -104.7067
24.2887 | -104.6922
24.2925 | -104.6779
24.2963 | -104.6635
24.3001 | -104.6487
24.3038 | -104.6339
24.3076 | -104.6206
24.3114 | -104.6041
24.3152 | -104.5893
24.3189 | -104.5732
24.3227 | -104.558
24.3265 | -104.5427
24.3303 | -104.5274
24.334 | -104.5132
24.3378 | -104.4975
24.3416 | -104.4798
24.3454 | -104.4616
24.3491 | -104.4434
24.3529 | -104.4267
24.3567 | -104.4089
24.3605 | -104.39
24.3642 | -104.3722
24.368 | -104.356
24.3718 | -104.3396
24.3756 | -104.3237
24.3793 | -104.3073
24.3831 | -104.2929
24.3869 | -104.2788
24.3907 | -104.2644
24.3944 | -104.2488
24.3982 | -104.2347
24.402 | -104.2215
24.4058 | -104.2103
24.4095 | -104.1995
24.4133 | -104.1892
24.4171 | -104.1792
24.4209 | -104.1708
24.4246 | -104.163
24.4284 | -104.1561
24.4322 | -104.1473
24.436 | -104.1403
24.4397 | -104.1354
24.4435 | -104.1323
24.4473 | -104.1288
24.451 | -104.1255
24.4548 | -104.1215
24.4586 | -104.1197
24.4624 | -104.1202
24.4661 | -104.1203
24.4699 | -104.1217
24.4737 | -104.1217
24.4775 | -104.1241
24.4812 | -104.1268
24.485 | -104.1311
24.4888 | -104.1352
24.4926 | -104.1395
24.4963 | -104.1439
24.5001 | -104.146
24.5039 | -104.1514
24.5077 | -104.156
24.5114 | -104.1611
24.5152 | -104.1663
24.519 | -104.1708
24.5228 | -104.1756
24.5265 | -104.1807
24.5303 | -104.1875
24.5341 | -104.1935
24.5379 | -104.199
24.5416 | -104.2054
24.5454 | -104.2118
24.5492 | -104.217
24.553 | -104.2193
24.5567 | -104.2233
24.5605 | -104.2284
24.5643 | -104.2318
24.5681 | -104.2355
24.5718 | -104.2387
24.5756 | -104.2416
24.5794 | -104.2429
24.5832 | -104.2432
24.5869 | -104.2452
24.5907 | -104.2453
24.5945 | -104.2456
24.5983 | -104.2462
24.602 | -104.2474
24.6058 | -104.2492
24.6096 | -104.2513
24.6134 | -104.2533
24.6171 | -104.2534
24.6209 | -104.2541
24.6247 | -104.2568
24.6285 | -104.2594
24.6322 | -104.2604
24.636 | -104.2624
24.6398 | -104.2643
24.6435 | -104.2643
24.6473 | -104.2641
24.6511 | -104.2644
24.6549 | -104.2635
24.6586 | -104.2627
24.6624 | -104.2586
24.6662 | -104.2556
24.67 | -104.253
24.6737 | -104.2506
24.6775 | -104.2486
24.6813 | -104.2477
24.6851 | -104.248
24.6888 | -104.249
24.6926 | -104.2499
24.6964 | -104.2504
24.7002 | -104.251
24.7039 | -104.2511
24.7077 | -104.2523
24.7115 | -104.2528
24.7153 | -104.2524
24.719 | -104.2525
24.7228 | -104.251
24.7266 | -104.2516
24.7304 | -104.2506
24.7341 | -104.2518
24.7379 | -104.2545
24.7417 | -104.2581
24.7455 | -104.261
24.7492 | -104.2648
24.753 | -104.2678
24.7568 | -104.2705
24.7606 | -104.2735
24.7643 | -104.2775
24.7681 | -104.2802
24.7719 | -104.2816
24.7757 | -104.2831
24.7794 | -104.2882
24.7832 | -104.2937
24.787 | -104.2984
24.7908 | -104.3031
24.7945 | -104.3095
24.7983 | -104.3147
24.8021 | -104.3219
24.8058 | -104.3284
24.8096 | -104.3346
24.8134 | -104.3413
24.8172 | -104.3464
24.8209 | -104.3532
24.8247 | -104.3587
24.8285 | -104.3651
24.8323 | -104.3734
24.836 | -104.3802
24.8398 | -104.3875
24.8436 | -104.3941
24.8474 | -104.4043
24.8511 | -104.4126
24.8549 | -104.4205
24.8587 | -104.4279
24.8625 | -104.4339
24.8662 | -104.4397
24.87 | -104.4461
24.8738 | -104.4532
24.8776 | -104.4607
24.8813 | -104.4676
24.8851 | -104.4737
24.8889 | -104.4803
24.8927 | -104.4867
24.8964 | -104.4921
24.9002 | -104.4963
24.904 | -104.5018
24.9078 | -104.5061
24.9115 | -104.5091
24.9153 | -104.5128
24.9191 | -104.5168
24.9229 | -104.5197
24.9266 | -104.5208
24.9304 | -104.5223
24.9342 | -104.524
24.938 | -104.5238
24.9417 | -104.5216
24.9455 | -104.5201
24.9493 | -104.5178
24.9531 | -104.5138
24.9568 | -104.5086
24.9606 | -104.5026
24.9644 | -104.4982
24.9682 | -104.4928
24.9719 | -104.4853
24.9757 | -104.4762
24.9795 | -104.4679
24.9833 | -104.4602
24.987 | -104.4525
24.9908 | -104.4437
24.9946 | -104.4368
24.9983 | -104.431
25.0021 | -104.4233
25.0059 | -104.4151
25.0097 | -104.4077
25.0134 | -104.4019
25.0172 | -104.3955
25.021 | -104.3881
25.0248 | -104.3806
25.0285 | -104.3737
25.0323 | -104.3635
25.0361 | -104.3536
25.0399 | -104.3446
25.0436 | -104.3348
25.0474 | -104.33
25.0512 | -104.3232
25.055 | -104.3163
25.0587 | -104.3079
25.0625 | -104.2992
25.0663 | -104.292
25.0701 | -104.2826
25.0738 | -104.2729
25.0776 | -104.2655
25.0814 | -104.2585
25.0852 | -104.2539
25.0889 | -104.2487
25.0927 | -104.2437
25.0965 | -104.2381
25.1003 | -104.2299
25.104 | -104.2211
25.1078 | -104.2133
25.1116 | -104.2061
25.1154 | -104.201
25.1191 | -104.1964
25.1229 | -104.1929
25.1267 | -104.1887
25.1305 | -104.1831
25.1342 | -104.1777
25.138 | -104.1721
25.1418 | -104.1668
25.1456 | -104.1626
25.1493 | -104.1568
25.1531 | -104.1516
25.1569 | -104.1464
25.1607 | -104.142
25.1644 | -104.1386
25.1682 | -104.134
25.172 | -104.1311
25.1757 | -104.1283
25.1795 | -104.1256
25.1833 | -104.1209
25.1871 | -104.1157
25.1908 | -104.1117
25.1946 | -104.1071
25.1984 | -104.1026
25.2022 | -104.0966
25.2059 | -104.0925
25.2097 | -104.088
25.2135 | -104.0826
25.2173 | -104.0794
25.221 | -104.0767
25.2248 | -104.0724
25.2286 | -104.0688
25.2324 | -104.0666
25.2361 | -104.0626
25.2399 | -104.0608
25.2437 | -104.0591
25.2475 | -104.0553
25.2512 | -104.0513
25.255 | -104.0495
25.2588 | -104.0475
25.2626 | -104.0466
25.2663 | -104.0459
25.2701 | -104.0457
25.2739 | -104.0429
25.2777 | -104.0417
25.2814 | -104.0423
25.2852 | -104.0443
25.289 | -104.0433
25.2928 | -104.0406
25.2965 | -104.0381
25.3003 | -104.0356
25.3041 | -104.0355
25.3079 | -104.0353
25.3116 | -104.0357
25.3154 | -104.0367
25.3192 | -104.0357
25.323 | -104.0364
25.3267 | -104.0365
25.3305 | -104.037
25.3343 | -104.0374
25.3381 | -104.0355
25.3418 | -104.0339
25.3456 | -104.0335
25.3494 | -104.0321
25.3531 | -104.0302
25.3569 | -104.0298
25.3607 | -104.0285
25.3645 | -104.0271
25.3682 | -104.0252
25.372 | -104.0228
25.3758 | -104.0206
25.3796 | -104.0191
25.3833 | -104.0164
25.3871 | -104.016
25.3909 | -104.0153
25.3947 | -104.0141
25.3984 | -104.0138
25.4022 | -104.0125
25.406 | -104.0106
25.4098 | -104.0083
25.4135 | -104.0062
25.4173 | -104.0072
25.4211 | -104.0096
25.4249 | -104.0111
25.4286 | -104.0124
25.4324 | -104.0157
25.4362 | -104.0172
25.44 | -104.0167
25.4437 | -104.0153
25.4475 | -104.0144
25.4513 | -104.0142
25.4551 | -104.0161
25.4588 | -104.0203
25.4626 | -104.0256
25.4664 | -104.032
25.4702 | -104.0377
25.4739 | -104.0403
25.4777 | -104.0432
25.4815 | -104.0476
25.4853 | -104.0542
25.489 | -104.0614
25.4928 | -104.0681
25.4966 | -104.0741
25.5004 | -104.079
25.5041 | -104.0832
25.5079 | -104.0898
25.5117 | -104.0965
25.5155 | -104.1036
25.5192 | -104.1105
25.523 | -104.1166
25.5268 | -104.1242
25.5306 | -104.1319
25.5343 | -104.1394
25.5381 | -104.1455
25.5419 | -104.152
25.5456 | -104.1565
25.5494 | -104.1623
25.5532 | -104.1678
25.557 | -104.1746
25.5607 | -104.181
25.5645 | -104.1855
25.5683 | -104.1889
25.5721 | -104.1944
25.5758 | -104.1992
25.5796 | -104.203
25.5834 | -104.208
25.5872 | -104.2119
25.5909 | -104.2162
25.5947 | -104.2193
25.5985 | -104.2248
25.6023 | -104.2286
25.606 | -104.2332
25.6098 | -104.2382
25.6136 | -104.2433
25.6174 | -104.2477
25.6211 | -104.2533
25.6249 | -104.2573
25.6287 | -104.2622
25.6325 | -104.2674
25.6362 | -104.2731
25.64 | -104.2789
25.6438 | -104.283
25.6476 | -104.2864
25.6513 | -104.2917
25.6551 | -104.299
25.6589 | -104.3074
25.6627 | -104.3157
25.6664 | -104.3203
25.6702 | -104.3252
25.674 | -104.3303
25.6778 | -104.335
25.6815 | -104.3395
25.6853 | -104.3444
25.6891 | -104.3514
25.6929 | -104.3568
25.6966 | -104.3586
25.7004 | -104.3639
25.7042 | -104.3686
25.708 | -104.3721
25.7117 | -104.3745
25.7155 | -104.3762
25.7193 | -104.3781
25.723 | -104.3818
25.7268 | -104.3858
25.7306 | -104.3891
25.7344 | -104.3912
25.7381 | -104.3927
25.7419 | -104.3925
25.7457 | -104.3919
25.7495 | -104.3919
25.7532 | -104.393
25.757 | -104.3943
25.7608 | -104.3946
25.7646 | -104.3941
25.7683 | -104.3931
25.7721 | -104.3917
25.7759 | -104.3899
25.7797 | -104.3877
25.7834 | -104.3831
25.7872 | -104.3797
25.791 | -104.3779
25.7948 | -104.3749
25.7985 | -104.372
25.8023 | -104.3708
25.8061 | -104.37
25.8099 | -104.3671
25.8136 | -104.3634
25.8174 | -104.3598
25.8212 | -104.3552
25.825 | -104.35
25.8287 | -104.344
25.8325 | -104.3395
25.8363 | -104.3363
25.8401 | -104.333
25.8438 | -104.3292
25.8476 | -104.3247
25.8514 | -104.3214
25.8552 | -104.3183
25.8589 | -104.315
25.8627 | -104.3126
25.8665 | -104.3115
25.8703 | -104.309
25.874 | -104.3062
25.8778 | -104.3046
25.8816 | -104.3036
25.8853 | -104.303
25.8891 | -104.3035
25.8929 | -104.3039
25.8967 | -104.3033
25.9004 | -104.305
25.9042 | -104.3074
25.908 | -104.3115
25.9118 | -104.3155
25.9155 | -104.3199
25.9193 | -104.3234
25.9231 | -104.3282
25.9269 | -104.3328
25.9306 | -104.3385
25.9344 | -104.3452
25.9382 | -104.3515
25.942 | -104.3583
25.9457 | -104.3648
25.9495 | -104.3719
25.9533 | -104.3765
25.9571 | -104.3818
25.9608 | -104.3877
25.9646 | -104.3944
25.9684 | -104.4028
25.9722 | -104.4123
25.9759 | -104.4228
25.9797 | -104.4338
25.9835 | -104.4429
25.9873 | -104.4497
25.991 | -104.4582
25.9948 | -104.4675
25.9986 | -104.4754
26.0024 | -104.4844
26.0061 | -104.4908
26.0099 | -104.5005
26.0137 | -104.5077
26.0175 | -104.5154
26.0212 | -104.523
26.025 | -104.5293
26.0288 | -104.5341
26.0326 | -104.5395
26.0363 | -104.5449
26.0401 | -104.5481
26.0439 | -104.5505
26.0477 | -104.5543
26.0514 | -104.5577
26.0552 | -104.5601
26.059 | -104.5621
26.0628 | -104.5661
26.0665 | -104.5682
26.0703 | -104.5679
26.0741 | -104.5666
26.0778 | -104.5637
26.0816 | -104.5624
26.0854 | -104.5613
26.0892 | -104.5589
26.0929 | -104.5552
26.0967 | -104.5511
26.1005 | -104.5463
26.1043 | -104.54
26.108 | -104.536
26.1118 | -104.5302
26.1156 | -104.5215
26.1194 | -104.513
26.1231 | -104.5039
26.1269 | -104.4964
26.1307 | -104.4881
26.1345 | -104.477
26.1382 | -104.47
26.142 | -104.4624
26.1458 | -104.4529
26.1496 | -104.4441
26.1533 | -104.4371
26.1571 | -104.4289
26.1609 | -104.4204
26.1647 | -104.4124
26.1684 | -104.404
26.1722 | -104.394
26.176 | -104.3855
26.1798 | -104.3774
26.1835 | -104.3704
26.1873 | -104.3626
26.1911 | -104.356
26.1949 | -104.3498
26.1986 | -104.3446
26.2024 | -104.3384
26.2062 | -104.3307
26.21 | -104.3237
26.2137 | -104.3176
26.2175 | -104.3104
26.2213 | -104.3043
26.2251 | -104.2977
26.2288 | -104.2923
26.2326 | -104.2879
26.2364 | -104.2817
26.2402 | -104.2766
26.2439 | -104.2723
26.2477 | -104.2674
26.2515 | -104.2632
26.2552 | -104.2605
26.259 | -104.257
26.2628 | -104.2506
26.2666 | -104.2459
26.2703 | -104.2418
26.2741 | -104.2384
26.2779 | -104.2343
26.2817 | -104.2307
26.2854 | -104.2272
26.2892 | -104.2248
26.293 | -104.2227
26.2968 | -104.2213
26.3005 | -104.2197
26.3043 | -104.2166
26.3081 | -104.2133
26.3119 | -104.2125
26.3156 | -104.212
26.3194 | -104.2109
26.3232 | -104.2129
26.327 | -104.2141
26.3307 | -104.2145
26.3345 | -104.2129
26.3383 | -104.2104
26.3421 | -104.2085
26.3458 | -104.2064
26.3496 | -104.2046
26.3534 | -104.2049
26.3572 | -104.2064
26.3609 | -104.2063
26.3647 | -104.2045
26.3685 | -104.2056
26.3723 | -104.2067
26.376 | -104.2076
26.3798 | -104.2071
26.3836 | -104.2095
26.3874 | -104.2117
26.3911 | -104.2115
26.3949 | -104.2125
26.3987 | -104.2143
26.4025 | -104.2159
26.4062 | -104.2201
26.41 | -104.2259
26.4138 | -104.2304
26.4176 | -104.2339
26.4213 | -104.2383
26.4251 | -104.2432
26.4289 | -104.248
26.4326 | -104.2529
26.4364 | -104.2582
26.4402 | -104.2648
26.444 | -104.2732
26.4477 | -104.2828
26.4515 | -104.2906
26.4553 | -104.2972
26.4591 | -104.3045
26.4628 | -104.3128
26.4666 | -104.3222
26.4704 | -104.3306
26.4742 | -104.3388
26.4779 | -104.347
26.4817 | -104.3571
26.4855 | -104.3668
26.4893 | -104.3742
26.493 | -104.3815
26.4968 | -104.3883
26.5006 | -104.395
26.5044 | -104.4035
26.5081 | -104.4099
26.5119 | -104.4154
26.5157 | -104.421
26.5195 | -104.4279
26.5232 | -104.4334
26.527 | -104.4386
26.5308 | -104.4422
26.5346 | -104.4461
26.5383 | -104.4502
26.5421 | -104.454
26.5459 | -104.4581
26.5497 | -104.4607
26.5534 | -104.4638
26.5572 | -104.4646
26.561 | -104.4652
26.5648 | -104.4665
26.5685 | -104.4681
26.5723 | -104.4685
26.5761 | -104.4693
26.5799 | -104.4706
26.5836 | -104.4732
26.5874 | -104.4765
26.5912 | -104.4792
26.595 | -104.4788
26.5987 | -104.4796
26.6025 | -104.4801
26.6063 | -104.4811
26.6101 | -104.4834
26.6138 | -104.4852
26.6176 | -104.4859
26.6214 | -104.486
26.6251 | -104.4846
26.6289 | -104.4846
26.6327 | -104.4856
26.6365 | -104.487
26.6402 | -104.4876
26.644 | -104.4875
26.6478 | -104.4864
26.6516 | -104.4862
26.6553 | -104.4857
26.6591 | -104.4864
26.6629 | -104.4872
26.6667 | -104.488
26.6704 | -104.4888
26.6742 | -104.4886
26.678 | -104.4901
26.6818 | -104.493
26.6855 | -104.4952
26.6893 | -104.496
26.6931 | -104.4965
26.6969 | -104.4984
26.7006 | -104.5004
26.7044 | -104.5032
26.7082 | -104.5049
26.712 | -104.509
26.7157 | -104.5142
26.7195 | -104.5172
26.7233 | -104.5212
26.7271 | -104.5246
26.7308 | -104.5284
26.7346 | -104.531
26.7384 | -104.5327
26.7422 | -104.5364
26.7459 | -104.5405
26.7497 | -104.5448
26.7535 | -104.5488
26.7573 | -104.5519
26.761 | -104.5531
26.7648 | -104.5552
26.7686 | -104.5589
26.7724 | -104.5599
26.7761 | -104.5626
26.7799 | -104.5655
26.7837 | -104.566
26.7874 | -104.5654
26.7912 | -104.5655
26.795 | -104.5661
26.7988 | -104.5651
26.8025 | -104.5637
26.8063 | -104.5627
26.8101 | -104.5616
26.8139 | -104.5611
26.8176 | -104.5605
26.8214 | -104.5575
26.8252 | -104.5551
26.829 | -104.5536
26.8327 | -104.5515
26.8365 | -104.549
26.8403 | -104.5472
26.8441 | -104.5446
26.8478 | -104.5427
26.8516 | -104.5403
26.8554 | -104.5354
26.8592 | -104.531
26.8629 | -104.5262
26.8667 | -104.5226
26.8705 | -104.5181
26.8743 | -104.5137
26.878 | -104.5103
26.8818 | -104.5056
26.8856 | -104.5018
26.8894 | -104.4987
26.8931 | -104.496
26.8969 | -104.492
26.9007 | -104.4894
26.9045 | -104.485
26.9082 | -104.4795
26.912 | -104.4749
26.9158 | -104.4716
26.9196 | -104.4691
26.9233 | -104.467
26.9271 | -104.4637
26.9309 | -104.46
26.9347 | -104.4568
26.9384 | -104.4546
26.9422 | -104.4513
26.946 | -104.4467
26.9498 | -104.4431
26.9535 | -104.4403
26.9573 | -104.4386
26.9611 | -104.4344
26.9648 | -104.429
26.9686 | -104.4234
26.9724 | -104.4199
26.9762 | -104.4161
26.9799 | -104.4112
26.9837 | -104.406
26.9875 | -104.4005
26.9913 | -104.3919
26.995 | -104.3833
26.9988 | -104.3754
27.0026 | -104.3704
27.0064 | -104.3614
27.0101 | -104.3539
27.0139 | -104.3484
27.0177 | -104.3414
27.0215 | -104.3334
27.0252 | -104.3249
27.029 | -104.317
27.0328 | -104.3092
27.0366 | -104.2994
27.0403 | -104.2895
27.0441 | -104.2795
27.0479 | -104.2694
27.0517 | -104.26
27.0554 | -104.2473
27.0592 | -104.2355
27.063 | -104.2233
27.0668 | -104.2117
27.0705 | -104.2006
27.0743 | -104.1889
27.0781 | -104.1765
27.0819 | -104.1637
27.0856 | -104.1508
27.0894 | -104.1395
27.0932 | -104.1272
27.097 | -104.1128
27.1007 | -104.0997
27.1045 | -104.0877
27.1083 | -104.074
27.1121 | -104.0596
27.1158 | -104.0482
27.1196 | -104.0373
27.1234 | -104.0246
27.1272 | -104.0117
27.1309 | -104.0003
27.1347 | -103.9905
27.1385 | -103.9809
27.1423 | -103.973
27.146 | -103.9638
27.1498 | -103.955
27.1536 | -103.9473
27.1573 | -103.9382
27.1611 | -103.9307
27.1649 | -103.9248
27.1687 | -103.9196
27.1724 | -103.9144
27.1762 | -103.9099
27.18 | -103.9083
27.1838 | -103.9068
27.1875 | -103.9046
27.1913 | -103.9032
27.1951 | -103.9025
27.1989 | -103.9003
27.2026 | -103.9014
27.2064 | -103.9033
27.2102 | -103.9061
27.214 | -103.9104
27.2177 | -103.9146
27.2215 | -103.9174
27.2253 | -103.929
27.2291 | -103.9352
27.2328 | -103.9414
27.2366 | -103.9484
27.2404 | -103.9581
27.2442 | -103.967
27.2479 | -103.9779
27.2517 | -103.9887
27.2555 | -103.9984
27.2593 | -104.0068
27.263 | -104.0162
27.2668 | -104.0269
27.2706 | -104.0376
27.2744 | -104.0482
27.2781 | -104.0601
27.2819 | -104.0711
27.2857 | -104.0818
27.2895 | -104.093
27.2932 | -104.1054
27.297 | -104.1183
27.3008 | -104.1297
27.3046 | -104.1392
27.3083 | -104.1485
27.3121 | -104.1579
27.3159 | -104.1666
27.3197 | -104.176
27.3234 | -104.184
27.3272 | -104.192
27.331 | -104.2005
27.3347 | -104.2086
27.3385 | -104.2151
27.3423 | -104.2201
27.3461 | -104.2247
27.3498 | -104.2286
27.3536 | -104.2332
27.3574 | -104.2367
27.3612 | -104.24
27.3649 | -104.2442
27.3687 | -104.2481
27.3725 | -104.2507
27.3763 | -104.2529
27.38 | -104.254
27.3838 | -104.2527
27.3876 | -104.2485
27.3914 | -104.247
27.3951 | -104.2426
27.3989 | -104.2376
27.4027 | -104.2343
27.4065 | -104.2314
27.4102 | -104.2258
27.414 | -104.2198
27.4178 | -104.2137
27.4216 | -104.2077
27.4253 | -104.2002
27.4291 | -104.1936
27.4329 | -104.188
27.4367 | -104.1816
27.4404 | -104.1755
27.4442 | -104.1687
27.448 | -104.1602
27.4518 | -104.1505
27.4555 | -104.1405
27.4593 | -104.132
27.4631 | -104.1234
27.4669 | -104.1155
27.4706 | -104.106
27.4744 | -104.0974
27.4782 | -104.0887
27.482 | -104.0821
27.4857 | -104.0757
27.4895 | -104.0676
27.4933 | -104.0617
27.4971 | -104.0547
27.5008 | -104.0461
27.5046 | -104.0379
27.5084 | -104.0328
27.5121 | -104.0272
27.5159 | -104.0199
27.5197 | -104.0159
27.5235 | -104.0115
27.5272 | -104.0076
27.531 | -104.006
27.5348 | -104.0047
27.5386 | -104.0046
27.5423 | -104.0029
27.5461 | -104.0006
27.5499 | -104.0007
27.5537 | -104.0028
27.5574 | -104.0058
27.5612 | -104.0101
27.565 | -104.0136
27.5688 | -104.0176
27.5725 | -104.0217
27.5763 | -104.027
27.5801 | -104.0313
27.5839 | -104.0371
27.5876 | -104.042
27.5914 | -104.0487
27.5952 | -104.0567
27.599 | -104.0644
27.6027 | -104.0737
27.6065 | -104.0843
27.6103 | -104.0944
27.6141 | -104.103
27.6178 | -104.1122
27.6216 | -104.1221
27.6254 | -104.1325
27.6292 | -104.14
27.6329 | -104.1472
27.6367 | -104.1536
27.6405 | -104.1604
27.6443 | -104.1691
27.648 | -104.1783
27.6518 | -104.1854
27.6556 | -104.1924
27.6594 | -104.199
27.6631 | -104.2048
27.6669 | -104.2102
27.6707 | -104.2166
27.6745 | -104.2213
27.6782 | -104.2258
27.682 | -104.2311
27.6858 | -104.2342
27.6895 | -104.2376
27.6933 | -104.2418
27.6971 | -104.2452
27.7009 | -104.2474
27.7046 | -104.2512
27.7084 | -104.2538
27.7122 | -104.2559
27.716 | -104.2567
27.7197 | -104.256
27.7235 | -104.2534
27.7273 | -104.251
27.7311 | -104.2494
27.7348 | -104.2467
27.7386 | -104.2428
27.7424 | -104.2401
27.7462 | -104.2369
27.7499 | -104.2313
27.7537 | -104.2254
27.7575 | -104.2194
27.7613 | -104.2139
27.765 | -104.2086
27.7688 | -104.1993
27.7726 | -104.1959
27.7764 | -104.193
27.7801 | -104.188
27.7839 | -104.1816
27.7877 | -104.176
27.7915 | -104.1693
27.7952 | -104.1589
27.799 | -104.149
27.8028 | -104.1428
27.8066 | -104.1372
27.8103 | -104.1297
27.8141 | -104.1226
27.8179 | -104.1155
27.8217 | -104.1091
27.8254 | -104.103
27.8292 | -104.0972
27.833 | -104.0918
27.8368 | -104.086
27.8405 | -104.082
27.8443 | -104.0789
27.8481 | -104.0749
27.8519 | -104.0689
27.8556 | -104.0647
27.8594 | -104.0614
27.8632 | -104.0569
27.8669 | -104.0513
27.8707 | -104.0474
27.8745 | -104.0434
27.8783 | -104.0383
27.882 | -104.0343
27.8858 | -104.0302
27.8896 | -104.0271
27.8934 | -104.0237
27.8971 | -104.0211
27.9009 | -104.0168
27.9047 | -104.0121
27.9085 | -104.0097
27.9122 | -104.0062
27.916 | -104.0038
27.9198 | -104.0006
27.9236 | -103.9945
27.9273 | -103.9877
27.9311 | -103.9814
27.9349 | -103.9766
27.9387 | -103.9738
27.9424 | -103.9707
27.9462 | -103.9659
27.95 | -103.9618
27.9538 | -103.9582
27.9575 | -103.953
27.9613 | -103.9476
27.9651 | -103.9463
27.9689 | -103.9435
27.9726 | -103.9415
27.9764 | -103.9416
27.9802 | -103.9402
27.984 | -103.937
27.9877 | -103.9341
27.9915 | -103.9313
27.9953 | -103.9284
27.9991 | -103.9268
28.0028 | -103.9246