## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 221 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 220
Hood | 02-92B/P | 158
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1015 | hPa |  | PRC PRO 01
Humidity | 40.9 | %/Hr
Internal temperature of calibration bench | 23.18 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1015 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.01 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.05 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3064 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 04/05/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 14/05/2018
Engineer | Anthony HUE | Approval |  | 18/05/2018
CEO | Robert GLEMOT | Approval |  | 18/05/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4719 | 57.8623
0.012209 | -36.0444 | 38.2438
0.018313 | -34.9117 | 27.9777
0.024417 | -34.4359 | 21.9032
0.030521 | -34.3637 | 18.297
0.036626 | -34.2224 | 15.212
0.04273 | -34.1278 | 12.7201
0.048834 | -34.0874 | 11.312
0.054938 | -34.0059 | 10.2663
0.061043 | -33.9297 | 9.1264
0.067147 | -33.9223 | 8.2383
0.073251 | -33.9325 | 7.4322
0.079355 | -33.9685 | 6.7494
0.08546 | -33.9965 | 6.1382
0.091564 | -33.9936 | 5.7232
0.097668 | -33.9687 | 5.557
0.10377 | -33.9385 | 5.3282
0.10988 | -33.9156 | 5.06
0.11598 | -33.9046 | 4.7644
0.12209 | -33.9092 | 4.4831
0.12819 | -33.9254 | 4.1939
0.13429 | -33.9425 | 3.8791
0.1404 | -33.9464 | 3.5492
0.1465 | -33.9484 | 3.2526
0.15261 | -33.9433 | 3.0032
0.15871 | -33.9329 | 2.8342
0.16482 | -33.9204 | 2.7478
0.17092 | -33.9169 | 2.7518
0.17702 | -33.9133 | 2.7062
0.18313 | -33.9113 | 2.6518
0.18923 | -33.9099 | 2.5478
0.19534 | -33.907 | 2.4126
0.20144 | -33.9044 | 2.2504
0.20754 | -33.9003 | 2.1143
0.21365 | -33.8989 | 1.9903
0.21975 | -33.8983 | 1.8903
0.22586 | -33.9004 | 1.8069
0.23196 | -33.9041 | 1.7468
0.23807 | -33.9101 | 1.7001
0.24417 | -33.9159 | 1.6509
0.25027 | -33.9215 | 1.5958
0.25638 | -33.9247 | 1.5367
0.26248 | -33.927 | 1.4655
0.26859 | -33.928 | 1.3905
0.27469 | -33.9283 | 1.3037
0.2808 | -33.9296 | 1.211
0.2869 | -33.9308 | 1.1247
0.293 | -33.9321 | 1.0441
0.29911 | -33.9336 | 0.9751
0.30521 | -33.9354 | 0.9096
0.31132 | -33.9374 | 0.87893
0.31742 | -33.94 | 0.83985
0.32353 | -33.9412 | 0.81343
0.32963 | -33.9433 | 0.79697
0.33573 | -33.9453 | 0.79262
0.34184 | -33.9479 | 0.78552
0.34794 | -33.951 | 0.77059
0.35405 | -33.9543 | 0.75629
0.36015 | -33.958 | 0.73944
0.36626 | -33.9609 | 0.71734
0.37236 | -33.964 | 0.68078
0.37846 | -33.9671 | 0.63244
0.38457 | -33.9681 | 0.56626
0.39067 | -33.9696 | 0.51587
0.39678 | -33.9702 | 0.45846
0.40288 | -33.9705 | 0.3973
0.40899 | -33.9701 | 0.3399
0.41509 | -33.9689 | 0.28533
0.42119 | -33.9686 | 0.24212
0.4273 | -33.9683 | 0.19396
0.4334 | -33.9679 | 0.14508
0.43951 | -33.9683 | 0.098688
0.44561 | -33.9694 | 0.057525
0.45172 | -33.9717 | 0.026519
0.45782 | -33.9728 | -0.011461
0.46392 | -33.9748 | -0.049663
0.47003 | -33.9763 | -0.081917
0.47613 | -33.9776 | -0.11788
0.48224 | -33.9784 | -0.15633
0.48834 | -33.9769 | -0.19877
0.49445 | -33.9765 | -0.23934
0.50055 | -33.9755 | -0.27626
0.50665 | -33.9742 | -0.31833
0.51276 | -33.9724 | -0.36085
0.51886 | -33.9705 | -0.40331
0.52497 | -33.9681 | -0.43297
0.53107 | -33.9665 | -0.47014
0.53717 | -33.965 | -0.50497
0.54328 | -33.963 | -0.5326
0.54938 | -33.9622 | -0.55977
0.55549 | -33.9617 | -0.58388
0.56159 | -33.9628 | -0.59259
0.5677 | -33.9627 | -0.60612
0.5738 | -33.9637 | -0.62211
0.5799 | -33.9645 | -0.63577
0.58601 | -33.9657 | -0.64106
0.59211 | -33.9678 | -0.63836
0.59822 | -33.9693 | -0.64169
0.60432 | -33.9711 | -0.64945
0.61043 | -33.9726 | -0.64957
0.61653 | -33.974 | -0.653
0.62263 | -33.976 | -0.65424
0.62874 | -33.9766 | -0.65685
0.63484 | -33.9776 | -0.6578
0.64095 | -33.978 | -0.66005
0.64705 | -33.9784 | -0.66222
0.65316 | -33.9784 | -0.66875
0.65926 | -33.9784 | -0.68134
0.66536 | -33.9767 | -0.69888
0.67147 | -33.976 | -0.7085
0.67757 | -33.9755 | -0.72021
0.68368 | -33.9744 | -0.73984
0.68978 | -33.9734 | -0.75719
0.69589 | -33.9722 | -0.77746
0.70199 | -33.9711 | -0.8057
0.70809 | -33.97 | -0.8267
0.7142 | -33.9691 | -0.8503
0.7203 | -33.9677 | -0.87261
0.72641 | -33.967 | -0.89983
0.73251 | -33.9663 | -0.92312
0.73862 | -33.9669 | -0.94873
0.74472 | -33.9666 | -0.97161
0.75082 | -33.9664 | -0.99771
0.75693 | -33.9671 | -1.0233
0.76303 | -33.968 | -1.0486
0.76914 | -33.9707 | -1.0693
0.77524 | -33.9722 | -1.0902
0.78135 | -33.9741 | -1.1116
0.78745 | -33.9761 | -1.1325
0.79355 | -33.9778 | -1.1529
0.79966 | -33.9802 | -1.1727
0.80576 | -33.9832 | -1.1921
0.81187 | -33.9853 | -1.2151
0.81797 | -33.988 | -1.2376
0.82408 | -33.9903 | -1.2589
0.83018 | -33.9926 | -1.28
0.83628 | -33.9952 | -1.2983
0.84239 | -33.9966 | -1.3205
0.84849 | -33.9985 | -1.3443
0.8546 | -34.0003 | -1.3649
0.8607 | -34.0019 | -1.3873
0.86681 | -34.0032 | -1.4073
0.87291 | -34.0045 | -1.4271
0.87901 | -34.005 | -1.4475
0.88512 | -34.0059 | -1.4669
0.89122 | -34.0068 | -1.4864
0.89733 | -34.0072 | -1.5026
0.90343 | -34.0077 | -1.5214
0.90953 | -34.008 | -1.5359
0.91564 | -34.0079 | -1.552
0.92174 | -34.0082 | -1.5671
0.92785 | -34.0083 | -1.5814
0.93395 | -34.0084 | -1.5946
0.94006 | -34.0087 | -1.6065
0.94616 | -34.0083 | -1.6164
0.95226 | -34.0085 | -1.6276
0.95837 | -34.0087 | -1.6405
0.96447 | -34.009 | -1.6518
0.97058 | -34.0088 | -1.664
0.97668 | -34.0085 | -1.6761
0.98279 | -34.0081 | -1.6867
0.98889 | -34.0078 | -1.6943
0.99499 | -34.0077 | -1.7064
1.0011 | -34.0076 | -1.7144
1.0072 | -34.0071 | -1.7235
1.0133 | -34.0066 | -1.731
1.0194 | -34.0059 | -1.7421
1.0255 | -34.0054 | -1.7542
1.0316 | -34.0045 | -1.7637
1.0377 | -34.0038 | -1.7757
1.0438 | -34.0028 | -1.7862
1.0499 | -34.0015 | -1.8009
1.056 | -34.0003 | -1.8141
1.0621 | -33.9993 | -1.8274
1.0682 | -33.9984 | -1.8378
1.0743 | -33.9973 | -1.8531
1.0805 | -33.9964 | -1.8667
1.0866 | -33.9947 | -1.8823
1.0927 | -33.9933 | -1.8962
1.0988 | -33.9922 | -1.911
1.1049 | -33.9908 | -1.9248
1.111 | -33.9894 | -1.9406
1.1171 | -33.9881 | -1.9549
1.1232 | -33.9872 | -1.9712
1.1293 | -33.9863 | -1.9851
1.1354 | -33.9857 | -2.0004
1.1415 | -33.9849 | -2.0168
1.1476 | -33.9838 | -2.0334
1.1537 | -33.9828 | -2.0488
1.1598 | -33.9828 | -2.0652
1.1659 | -33.9822 | -2.0788
1.172 | -33.9816 | -2.0932
1.1781 | -33.981 | -2.1076
1.1842 | -33.9805 | -2.1224
1.1903 | -33.9809 | -2.1328
1.1964 | -33.9805 | -2.1468
1.2025 | -33.9804 | -2.1615
1.2086 | -33.9807 | -2.174
1.2147 | -33.9805 | -2.1867
1.2209 | -33.9808 | -2.1973
1.227 | -33.9818 | -2.2071
1.2331 | -33.9823 | -2.2191
1.2392 | -33.9826 | -2.2308
1.2453 | -33.9831 | -2.2413
1.2514 | -33.9836 | -2.25
1.2575 | -33.9843 | -2.2601
1.2636 | -33.9855 | -2.2671
1.2697 | -33.9863 | -2.2748
1.2758 | -33.9872 | -2.2854
1.2819 | -33.9881 | -2.294
1.288 | -33.989 | -2.304
1.2941 | -33.9901 | -2.3106
1.3002 | -33.9911 | -2.3164
1.3063 | -33.9918 | -2.3235
1.3124 | -33.9927 | -2.3308
1.3185 | -33.9937 | -2.3382
1.3246 | -33.9947 | -2.3454
1.3307 | -33.9959 | -2.3521
1.3368 | -33.997 | -2.3589
1.3429 | -33.9979 | -2.3661
1.349 | -33.9986 | -2.3739
1.3551 | -33.9996 | -2.3815
1.3613 | -34.0003 | -2.3875
1.3674 | -34.0012 | -2.3965
1.3735 | -34.0018 | -2.4048
1.3796 | -34.0024 | -2.4122
1.3857 | -34.0028 | -2.4197
1.3918 | -34.0035 | -2.4286
1.3979 | -34.0043 | -2.4362
1.404 | -34.0049 | -2.4474
1.4101 | -34.0055 | -2.4568
1.4162 | -34.006 | -2.4666
1.4223 | -34.0066 | -2.4777
1.4284 | -34.0072 | -2.4866
1.4345 | -34.0075 | -2.4976
1.4406 | -34.0078 | -2.512
1.4467 | -34.0084 | -2.5236
1.4528 | -34.009 | -2.5346
1.4589 | -34.0096 | -2.5464
1.465 | -34.0098 | -2.5566
1.4711 | -34.0105 | -2.5676
1.4772 | -34.0111 | -2.5819
1.4833 | -34.0117 | -2.5955
1.4894 | -34.0125 | -2.6078
1.4955 | -34.0132 | -2.6198
1.5016 | -34.0139 | -2.6326
1.5078 | -34.0143 | -2.6475
1.5139 | -34.0149 | -2.6611
1.52 | -34.0154 | -2.674
1.5261 | -34.0158 | -2.6868
1.5322 | -34.0162 | -2.702
1.5383 | -34.0167 | -2.7148
1.5444 | -34.0174 | -2.7277
1.5505 | -34.018 | -2.7421
1.5566 | -34.0187 | -2.7556
1.5627 | -34.0193 | -2.7703
1.5688 | -34.0201 | -2.7848
1.5749 | -34.0205 | -2.7978
1.581 | -34.0209 | -2.8119
1.5871 | -34.0214 | -2.8266
1.5932 | -34.0218 | -2.84
1.5993 | -34.0223 | -2.8554
1.6054 | -34.0228 | -2.8694
1.6115 | -34.0233 | -2.8848
1.6176 | -34.0236 | -2.8975
1.6237 | -34.0239 | -2.912
1.6298 | -34.0242 | -2.9248
1.6359 | -34.0246 | -2.9379
1.642 | -34.0251 | -2.9497
1.6482 | -34.0253 | -2.963
1.6543 | -34.0258 | -2.9774
1.6604 | -34.0262 | -2.9892
1.6665 | -34.0262 | -3.0017
1.6726 | -34.0264 | -3.0143
1.6787 | -34.0269 | -3.0271
1.6848 | -34.0266 | -3.0397
1.6909 | -34.0269 | -3.0528
1.697 | -34.0271 | -3.065
1.7031 | -34.027 | -3.0776
1.7092 | -34.0269 | -3.0895
1.7153 | -34.027 | -3.1011
1.7214 | -34.0266 | -3.1131
1.7275 | -34.0264 | -3.1256
1.7336 | -34.0262 | -3.1387
1.7397 | -34.0263 | -3.1508
1.7458 | -34.0262 | -3.1627
1.7519 | -34.0261 | -3.1745
1.758 | -34.026 | -3.1858
1.7641 | -34.0262 | -3.1968
1.7702 | -34.0259 | -3.2098
1.7763 | -34.0258 | -3.2205
1.7824 | -34.0256 | -3.2326
1.7885 | -34.0255 | -3.2448
1.7947 | -34.0255 | -3.2562
1.8008 | -34.0254 | -3.2667
1.8069 | -34.0253 | -3.2775
1.813 | -34.0253 | -3.2895
1.8191 | -34.0252 | -3.3021
1.8252 | -34.0252 | -3.3148
1.8313 | -34.0253 | -3.3268
1.8374 | -34.0254 | -3.3382
1.8435 | -34.0254 | -3.3495
1.8496 | -34.0253 | -3.3626
1.8557 | -34.0255 | -3.3753
1.8618 | -34.0258 | -3.3873
1.8679 | -34.0258 | -3.3991
1.874 | -34.0259 | -3.4107
1.8801 | -34.026 | -3.4231
1.8862 | -34.0263 | -3.4347
1.8923 | -34.0264 | -3.4467
1.8984 | -34.0265 | -3.4586
1.9045 | -34.0265 | -3.4697
1.9106 | -34.0265 | -3.4818
1.9167 | -34.0265 | -3.4938
1.9228 | -34.0266 | -3.5054
1.9289 | -34.0267 | -3.5172
1.9351 | -34.0268 | -3.5273
1.9412 | -34.0271 | -3.538
1.9473 | -34.0275 | -3.5501
1.9534 | -34.0277 | -3.5612
1.9595 | -34.0279 | -3.5729
1.9656 | -34.0282 | -3.5856
1.9717 | -34.0284 | -3.5971
1.9778 | -34.0288 | -3.6083
1.9839 | -34.0291 | -3.6196
1.99 | -34.0293 | -3.6303
1.9961 | -34.0294 | -3.6425
2.0022 | -34.0297 | -3.6541
2.0083 | -34.0297 | -3.665
2.0144 | -34.03 | -3.676
2.0205 | -34.0304 | -3.688
2.0266 | -34.0308 | -3.6991
2.0327 | -34.0311 | -3.7107
2.0388 | -34.0316 | -3.7226
2.0449 | -34.0318 | -3.7338
2.051 | -34.0322 | -3.7453
2.0571 | -34.0326 | -3.7566
2.0632 | -34.0328 | -3.7681
2.0693 | -34.033 | -3.7808
2.0754 | -34.0333 | -3.7922
2.0816 | -34.0337 | -3.8049
2.0877 | -34.0339 | -3.8165
2.0938 | -34.0344 | -3.8292
2.0999 | -34.0345 | -3.8411
2.106 | -34.0347 | -3.8529
2.1121 | -34.0349 | -3.8653
2.1182 | -34.0352 | -3.8758
2.1243 | -34.0356 | -3.8863
2.1304 | -34.0358 | -3.8974
2.1365 | -34.0361 | -3.9095
2.1426 | -34.0364 | -3.922
2.1487 | -34.0367 | -3.9334
2.1548 | -34.0373 | -3.9456
2.1609 | -34.0378 | -3.9572
2.167 | -34.0383 | -3.9685
2.1731 | -34.0387 | -3.9809
2.1792 | -34.0389 | -3.992
2.1853 | -34.0394 | -4.0044
2.1914 | -34.0397 | -4.0159
2.1975 | -34.0402 | -4.0276
2.2036 | -34.0405 | -4.0402
2.2097 | -34.0408 | -4.0519
2.2158 | -34.0412 | -4.0622
2.222 | -34.0415 | -4.0723
2.2281 | -34.0418 | -4.0839
2.2342 | -34.0422 | -4.0954
2.2403 | -34.0424 | -4.1058
2.2464 | -34.0427 | -4.1163
2.2525 | -34.043 | -4.1278
2.2586 | -34.0434 | -4.1382
2.2647 | -34.0437 | -4.1482
2.2708 | -34.0441 | -4.1595
2.2769 | -34.0445 | -4.1699
2.283 | -34.0447 | -4.181
2.2891 | -34.0449 | -4.1918
2.2952 | -34.0451 | -4.2021
2.3013 | -34.0453 | -4.2127
2.3074 | -34.0456 | -4.2228
2.3135 | -34.0458 | -4.2344
2.3196 | -34.046 | -4.2441
2.3257 | -34.0462 | -4.2545
2.3318 | -34.0463 | -4.2656
2.3379 | -34.0465 | -4.2761
2.344 | -34.0466 | -4.2875
2.3501 | -34.0468 | -4.2979
2.3562 | -34.0469 | -4.3077
2.3623 | -34.047 | -4.3177
2.3685 | -34.0471 | -4.3279
2.3746 | -34.0471 | -4.3392
2.3807 | -34.0474 | -4.3493
2.3868 | -34.0475 | -4.3598
2.3929 | -34.0477 | -4.3701
2.399 | -34.0477 | -4.3801
2.4051 | -34.048 | -4.3902
2.4112 | -34.0482 | -4.4003
2.4173 | -34.0483 | -4.4107
2.4234 | -34.0483 | -4.4212
2.4295 | -34.0485 | -4.4318
2.4356 | -34.0486 | -4.4425
2.4417 | -34.0487 | -4.4535
2.4478 | -34.0487 | -4.4643
2.4539 | -34.0489 | -4.4739
2.46 | -34.0489 | -4.4855
2.4661 | -34.049 | -4.4964
2.4722 | -34.0492 | -4.506
2.4783 | -34.0492 | -4.5159
2.4844 | -34.0492 | -4.5268
2.4905 | -34.0492 | -4.5376
2.4966 | -34.0493 | -4.548
2.5027 | -34.0494 | -4.5581
2.5089 | -34.0495 | -4.5686
2.515 | -34.0494 | -4.5791
2.5211 | -34.0494 | -4.5895
2.5272 | -34.0494 | -4.6001
2.5333 | -34.0495 | -4.6107
2.5394 | -34.0496 | -4.6211
2.5455 | -34.0496 | -4.6324
2.5516 | -34.0497 | -4.6422
2.5577 | -34.0499 | -4.6536
2.5638 | -34.05 | -4.6655
2.5699 | -34.05 | -4.6768
2.576 | -34.0499 | -4.6885
2.5821 | -34.05 | -4.6996
2.5882 | -34.05 | -4.7106
2.5943 | -34.0501 | -4.721
2.6004 | -34.0503 | -4.7316
2.6065 | -34.0503 | -4.7415
2.6126 | -34.0504 | -4.7514
2.6187 | -34.0505 | -4.7613
2.6248 | -34.0506 | -4.7717
2.6309 | -34.0507 | -4.7831
2.637 | -34.051 | -4.7948
2.6431 | -34.0511 | -4.8048
2.6492 | -34.0511 | -4.816
2.6554 | -34.0513 | -4.8266
2.6615 | -34.0513 | -4.8372
2.6676 | -34.0516 | -4.8486
2.6737 | -34.0519 | -4.8595
2.6798 | -34.0522 | -4.8704
2.6859 | -34.0523 | -4.8815
2.692 | -34.0525 | -4.8919
2.6981 | -34.0527 | -4.9028
2.7042 | -34.0529 | -4.9137
2.7103 | -34.053 | -4.9238
2.7164 | -34.0534 | -4.9346
2.7225 | -34.0536 | -4.9458
2.7286 | -34.0538 | -4.9565
2.7347 | -34.054 | -4.9665
2.7408 | -34.0543 | -4.9758
2.7469 | -34.0545 | -4.9869
2.753 | -34.0549 | -4.9969
2.7591 | -34.0552 | -5.0082
2.7652 | -34.0554 | -5.0197
2.7713 | -34.0556 | -5.03
2.7774 | -34.0559 | -5.0407
2.7835 | -34.056 | -5.0513
2.7896 | -34.0562 | -5.0612
2.7958 | -34.0565 | -5.0715
2.8019 | -34.0568 | -5.0813
2.808 | -34.0571 | -5.0917
2.8141 | -34.0575 | -5.1027
2.8202 | -34.0579 | -5.1133
2.8263 | -34.0582 | -5.1238
2.8324 | -34.0587 | -5.1352
2.8385 | -34.0592 | -5.1457
2.8446 | -34.0596 | -5.1561
2.8507 | -34.06 | -5.1666
2.8568 | -34.0604 | -5.1766
2.8629 | -34.0607 | -5.1867
2.869 | -34.0611 | -5.1976
2.8751 | -34.0615 | -5.2078
2.8812 | -34.062 | -5.219
2.8873 | -34.0624 | -5.2295
2.8934 | -34.0627 | -5.2403
2.8995 | -34.0631 | -5.2512
2.9056 | -34.0634 | -5.2621
2.9117 | -34.0639 | -5.2729
2.9178 | -34.0643 | -5.2833
2.9239 | -34.0648 | -5.2949
2.93 | -34.0653 | -5.3063
2.9361 | -34.0656 | -5.3168
2.9423 | -34.0661 | -5.327
2.9484 | -34.0666 | -5.3379
2.9545 | -34.0671 | -5.3487
2.9606 | -34.0676 | -5.3601
2.9667 | -34.068 | -5.3712
2.9728 | -34.0684 | -5.3823
2.9789 | -34.069 | -5.3932
2.985 | -34.0695 | -5.4045
2.9911 | -34.07 | -5.4158
2.9972 | -34.0706 | -5.4268
3.0033 | -34.0711 | -5.4384
3.0094 | -34.0715 | -5.4494
3.0155 | -34.0721 | -5.46
3.0216 | -34.0726 | -5.4713
3.0277 | -34.0729 | -5.483
3.0338 | -34.0734 | -5.4944
3.0399 | -34.0739 | -5.5058
3.046 | -34.0743 | -5.5174
3.0521 | -34.0747 | -5.5283
3.0582 | -34.075 | -5.5407
3.0643 | -34.0754 | -5.5522
3.0704 | -34.0758 | -5.5642
3.0765 | -34.0761 | -5.5752
3.0827 | -34.0765 | -5.5864
3.0888 | -34.0769 | -5.5977
3.0949 | -34.0774 | -5.6092
3.101 | -34.0778 | -5.6207
3.1071 | -34.0782 | -5.6328
3.1132 | -34.0787 | -5.6448
3.1193 | -34.0791 | -5.6564
3.1254 | -34.0796 | -5.6678
3.1315 | -34.08 | -5.6801
3.1376 | -34.0803 | -5.6916
3.1437 | -34.0808 | -5.7033
3.1498 | -34.0812 | -5.7158
3.1559 | -34.0816 | -5.7274
3.162 | -34.0818 | -5.7401
3.1681 | -34.0821 | -5.7524
3.1742 | -34.0825 | -5.7641
3.1803 | -34.0829 | -5.7759
3.1864 | -34.0833 | -5.7873
3.1925 | -34.0837 | -5.799
3.1986 | -34.084 | -5.8115
3.2047 | -34.0843 | -5.8241
3.2108 | -34.0845 | -5.8351
3.2169 | -34.0849 | -5.8472
3.223 | -34.0852 | -5.8591
3.2292 | -34.0854 | -5.8713
3.2353 | -34.0856 | -5.8821
3.2414 | -34.0859 | -5.8937
3.2475 | -34.0861 | -5.9052
3.2536 | -34.0865 | -5.917
3.2597 | -34.0867 | -5.9278
3.2658 | -34.0869 | -5.9387
3.2719 | -34.0872 | -5.9498
3.278 | -34.0875 | -5.9616
3.2841 | -34.0877 | -5.9727
3.2902 | -34.088 | -5.9844
3.2963 | -34.0883 | -5.9961
3.3024 | -34.0883 | -6.0073
3.3085 | -34.0887 | -6.0193
3.3146 | -34.0888 | -6.0304
3.3207 | -34.0891 | -6.0422
3.3268 | -34.0893 | -6.0539
3.3329 | -34.0895 | -6.0645
3.339 | -34.0896 | -6.0753
3.3451 | -34.0898 | -6.0867
3.3512 | -34.0899 | -6.0973
3.3573 | -34.0902 | -6.1086
3.3634 | -34.0904 | -6.1199
3.3696 | -34.0905 | -6.1314
3.3757 | -34.0906 | -6.1421
3.3818 | -34.0909 | -6.1529
3.3879 | -34.091 | -6.164
3.394 | -34.0911 | -6.1756
3.4001 | -34.0913 | -6.1864
3.4062 | -34.0913 | -6.1973
3.4123 | -34.0915 | -6.2077
3.4184 | -34.0917 | -6.2192
3.4245 | -34.0919 | -6.23
3.4306 | -34.0919 | -6.2405
3.4367 | -34.0921 | -6.2509
3.4428 | -34.0922 | -6.261
3.4489 | -34.0923 | -6.272
3.455 | -34.0924 | -6.2828
3.4611 | -34.0926 | -6.2941
3.4672 | -34.0927 | -6.3053
3.4733 | -34.0929 | -6.3164
3.4794 | -34.093 | -6.3271
3.4855 | -34.0931 | -6.3372
3.4916 | -34.0933 | -6.3473
3.4977 | -34.0934 | -6.3588
3.5038 | -34.0935 | -6.3696
3.5099 | -34.0936 | -6.3795
3.5161 | -34.0938 | -6.3897
3.5222 | -34.0938 | -6.3998
3.5283 | -34.0939 | -6.4105
3.5344 | -34.0939 | -6.4212
3.5405 | -34.0941 | -6.4316
3.5466 | -34.0943 | -6.4425
3.5527 | -34.0943 | -6.4532
3.5588 | -34.0943 | -6.4637
3.5649 | -34.0943 | -6.4741
3.571 | -34.0944 | -6.4845
3.5771 | -34.0945 | -6.4957
3.5832 | -34.0946 | -6.5051
3.5893 | -34.0948 | -6.5156
3.5954 | -34.095 | -6.5264
3.6015 | -34.0951 | -6.5373
3.6076 | -34.0953 | -6.5477
3.6137 | -34.0954 | -6.5574
3.6198 | -34.0957 | -6.5673
3.6259 | -34.0958 | -6.5782
3.632 | -34.096 | -6.5887
3.6381 | -34.0961 | -6.5994
3.6442 | -34.0962 | -6.6098
3.6503 | -34.0965 | -6.6196
3.6565 | -34.0967 | -6.6303
3.6626 | -34.0967 | -6.6406
3.6687 | -34.097 | -6.6512
3.6748 | -34.0972 | -6.6614
3.6809 | -34.0975 | -6.6719
3.687 | -34.0976 | -6.6828
3.6931 | -34.098 | -6.6928
3.6992 | -34.0981 | -6.7033
3.7053 | -34.0983 | -6.7136
3.7114 | -34.0984 | -6.7241
3.7175 | -34.0987 | -6.7347
3.7236 | -34.099 | -6.7443
3.7297 | -34.0992 | -6.755
3.7358 | -34.0994 | -6.766
3.7419 | -34.0995 | -6.7768
3.748 | -34.0998 | -6.7873
3.7541 | -34.0999 | -6.7977
3.7602 | -34.1003 | -6.808
3.7663 | -34.1005 | -6.8182
3.7724 | -34.1007 | -6.8287
3.7785 | -34.1011 | -6.839
3.7846 | -34.1012 | -6.8491
3.7907 | -34.1016 | -6.8602
3.7969 | -34.1018 | -6.871
3.803 | -34.1021 | -6.8808
3.8091 | -34.1023 | -6.8912
3.8152 | -34.1027 | -6.9016
3.8213 | -34.1029 | -6.9125
3.8274 | -34.1032 | -6.9225
3.8335 | -34.1036 | -6.9333
3.8396 | -34.1039 | -6.9439
3.8457 | -34.1043 | -6.9546
3.8518 | -34.1047 | -6.965
3.8579 | -34.1051 | -6.9754
3.864 | -34.1054 | -6.9863
3.8701 | -34.1057 | -6.9975
3.8762 | -34.106 | -7.0077
3.8823 | -34.1062 | -7.0185
3.8884 | -34.1066 | -7.0291
3.8945 | -34.1071 | -7.0401
3.9006 | -34.1074 | -7.0513
3.9067 | -34.1077 | -7.0625
3.9128 | -34.108 | -7.0734
3.9189 | -34.1084 | -7.0836
3.925 | -34.1086 | -7.0945
3.9311 | -34.1089 | -7.1043
3.9372 | -34.1093 | -7.1153
3.9434 | -34.1096 | -7.1259
3.9495 | -34.1101 | -7.1368
3.9556 | -34.1105 | -7.1471
3.9617 | -34.1108 | -7.1572
3.9678 | -34.1112 | -7.1674
3.9739 | -34.1115 | -7.1781
3.98 | -34.1118 | -7.1889
3.9861 | -34.1121 | -7.1995
3.9922 | -34.1123 | -7.2102
3.9983 | -34.1127 | -7.2202
4.0044 | -34.113 | -7.2301
4.0105 | -34.1133 | -7.2408
4.0166 | -34.1137 | -7.2516
4.0227 | -34.1141 | -7.2622
4.0288 | -34.1143 | -7.2735
4.0349 | -34.1145 | -7.2837
4.041 | -34.1148 | -7.2943
4.0471 | -34.1151 | -7.3049
4.0532 | -34.1154 | -7.3151
4.0593 | -34.1158 | -7.326
4.0654 | -34.1161 | -7.3363
4.0715 | -34.1164 | -7.3475
4.0776 | -34.1168 | -7.3581
4.0838 | -34.1171 | -7.3693
4.0899 | -34.1175 | -7.3796
4.096 | -34.1178 | -7.3902
4.1021 | -34.118 | -7.4004
4.1082 | -34.1183 | -7.4107
4.1143 | -34.1186 | -7.4207
4.1204 | -34.1189 | -7.4309
4.1265 | -34.1192 | -7.441
4.1326 | -34.1195 | -7.4511
4.1387 | -34.1199 | -7.4615
4.1448 | -34.1202 | -7.4721
4.1509 | -34.1204 | -7.4825
4.157 | -34.1207 | -7.4936
4.1631 | -34.121 | -7.5045
4.1692 | -34.1213 | -7.5148
4.1753 | -34.1216 | -7.5252
4.1814 | -34.1217 | -7.5351
4.1875 | -34.1221 | -7.546
4.1936 | -34.1224 | -7.5566
4.1997 | -34.1226 | -7.5665
4.2058 | -34.1228 | -7.5773
4.2119 | -34.1231 | -7.5875
4.218 | -34.1233 | -7.5975
4.2241 | -34.1235 | -7.6075
4.2303 | -34.1239 | -7.6177
4.2364 | -34.1241 | -7.6281
4.2425 | -34.1243 | -7.6385
4.2486 | -34.1245 | -7.6492
4.2547 | -34.1246 | -7.6591
4.2608 | -34.1248 | -7.6696
4.2669 | -34.1249 | -7.6794
4.273 | -34.1252 | -7.6899
4.2791 | -34.1254 | -7.7002
4.2852 | -34.1257 | -7.7109
4.2913 | -34.1258 | -7.7206
4.2974 | -34.126 | -7.7314
4.3035 | -34.1262 | -7.7422
4.3096 | -34.1264 | -7.7528
4.3157 | -34.1266 | -7.7629
4.3218 | -34.1268 | -7.7729
4.3279 | -34.127 | -7.7832
4.334 | -34.1272 | -7.7934
4.3401 | -34.1274 | -7.8031
4.3462 | -34.1277 | -7.8136
4.3523 | -34.128 | -7.8232
4.3584 | -34.1281 | -7.8328
4.3645 | -34.1283 | -7.8429
4.3707 | -34.1285 | -7.8526
4.3768 | -34.1287 | -7.8627
4.3829 | -34.1289 | -7.8734
4.389 | -34.1291 | -7.8835
4.3951 | -34.1293 | -7.8933
4.4012 | -34.1294 | -7.9039
4.4073 | -34.1296 | -7.9138
4.4134 | -34.1298 | -7.9237
4.4195 | -34.13 | -7.9338
4.4256 | -34.1303 | -7.9437
4.4317 | -34.1305 | -7.9537
4.4378 | -34.1307 | -7.9632
4.4439 | -34.1308 | -7.9728
4.45 | -34.1309 | -7.9829
4.4561 | -34.131 | -7.9929
4.4622 | -34.1312 | -8.003
4.4683 | -34.1314 | -8.013
4.4744 | -34.1316 | -8.0234
4.4805 | -34.1318 | -8.0328
4.4866 | -34.132 | -8.0428
4.4927 | -34.1322 | -8.0532
4.4988 | -34.1323 | -8.0631
4.5049 | -34.1326 | -8.0725
4.511 | -34.1328 | -8.082
4.5172 | -34.133 | -8.0919
4.5233 | -34.1331 | -8.1025
4.5294 | -34.1334 | -8.1122
4.5355 | -34.1335 | -8.1222
4.5416 | -34.1337 | -8.1321
4.5477 | -34.1339 | -8.1421
4.5538 | -34.1341 | -8.1514
4.5599 | -34.1344 | -8.1612
4.566 | -34.1345 | -8.1711
4.5721 | -34.1347 | -8.1811
4.5782 | -34.135 | -8.1914
4.5843 | -34.1352 | -8.2014
4.5904 | -34.1355 | -8.2114
4.5965 | -34.1357 | -8.2212
4.6026 | -34.1359 | -8.2318
4.6087 | -34.136 | -8.2416
4.6148 | -34.1362 | -8.252
4.6209 | -34.1365 | -8.2618
4.627 | -34.1367 | -8.2721
4.6331 | -34.1368 | -8.2818
4.6392 | -34.137 | -8.2919
4.6453 | -34.1372 | -8.3018
4.6514 | -34.1375 | -8.3118
4.6576 | -34.1376 | -8.3224
4.6637 | -34.1378 | -8.3318
4.6698 | -34.1379 | -8.3414
4.6759 | -34.1382 | -8.352
4.682 | -34.1383 | -8.3618
4.6881 | -34.1384 | -8.3717
4.6942 | -34.1387 | -8.3817
4.7003 | -34.1389 | -8.3916
4.7064 | -34.1391 | -8.4012
4.7125 | -34.1393 | -8.4117
4.7186 | -34.1395 | -8.4212
4.7247 | -34.1398 | -8.4314
4.7308 | -34.14 | -8.4412
4.7369 | -34.1402 | -8.4514
4.743 | -34.1404 | -8.4619
4.7491 | -34.1406 | -8.4724
4.7552 | -34.1408 | -8.4826
4.7613 | -34.1409 | -8.4928
4.7674 | -34.1411 | -8.5025
4.7735 | -34.1413 | -8.5132
4.7796 | -34.1416 | -8.5237
4.7857 | -34.1418 | -8.534
4.7918 | -34.142 | -8.5438
4.7979 | -34.1423 | -8.5537
4.8041 | -34.1425 | -8.5638
4.8102 | -34.1427 | -8.5734
4.8163 | -34.143 | -8.5834
4.8224 | -34.1432 | -8.5935
4.8285 | -34.1434 | -8.6031
4.8346 | -34.1437 | -8.6128
4.8407 | -34.1438 | -8.6229
4.8468 | -34.1441 | -8.6334
4.8529 | -34.1444 | -8.6439
4.859 | -34.1446 | -8.6544
4.8651 | -34.1448 | -8.6642
4.8712 | -34.145 | -8.6739
4.8773 | -34.1452 | -8.684
4.8834 | -34.1454 | -8.6942
4.8895 | -34.1456 | -8.704
4.8956 | -34.1458 | -8.714
4.9017 | -34.146 | -8.7238
4.9078 | -34.1462 | -8.7343
4.9139 | -34.1464 | -8.7438
4.92 | -34.1468 | -8.7535
4.9261 | -34.147 | -8.7631
4.9322 | -34.1473 | -8.7732
4.9383 | -34.1476 | -8.7832
4.9445 | -34.1478 | -8.7933
4.9506 | -34.148 | -8.8031
4.9567 | -34.1482 | -8.8128
4.9628 | -34.1484 | -8.8228
4.9689 | -34.1487 | -8.8335
4.975 | -34.1488 | -8.8439
4.9811 | -34.1491 | -8.8536
4.9872 | -34.1494 | -8.8635
4.9933 | -34.1498 | -8.8739
4.9994 | -34.15 | -8.8846
5.0055 | -34.1503 | -8.8947
5.0116 | -34.1505 | -8.9054
5.0177 | -34.1509 | -8.916
5.0238 | -34.1511 | -8.9266
5.0299 | -34.1515 | -8.9373
5.036 | -34.1518 | -8.9474
5.0421 | -34.152 | -8.9573
5.0482 | -34.1522 | -8.9675
5.0543 | -34.1525 | -8.9773
5.0604 | -34.1528 | -8.9874
5.0665 | -34.1531 | -8.9972
5.0726 | -34.1534 | -9.0079
5.0787 | -34.1536 | -9.0184
5.0848 | -34.1539 | -9.0287
5.091 | -34.1542 | -9.0386
5.0971 | -34.1545 | -9.0483
5.1032 | -34.1548 | -9.0589
5.1093 | -34.1551 | -9.0691
5.1154 | -34.1553 | -9.0789
5.1215 | -34.1555 | -9.0887
5.1276 | -34.1558 | -9.0985
5.1337 | -34.1562 | -9.1083
5.1398 | -34.1564 | -9.118
5.1459 | -34.1567 | -9.1286
5.152 | -34.1569 | -9.139
5.1581 | -34.1572 | -9.1495
5.1642 | -34.1575 | -9.1598
5.1703 | -34.1577 | -9.1699
5.1764 | -34.1581 | -9.1798
5.1825 | -34.1584 | -9.1896
5.1886 | -34.1587 | -9.1992
5.1947 | -34.1589 | -9.2094
5.2008 | -34.1592 | -9.2197
5.2069 | -34.1595 | -9.2302
5.213 | -34.1598 | -9.2401
5.2191 | -34.1601 | -9.2504
5.2252 | -34.1603 | -9.2606
5.2314 | -34.1606 | -9.2711
5.2375 | -34.1608 | -9.2809
5.2436 | -34.1611 | -9.2911
5.2497 | -34.1614 | -9.3009
5.2558 | -34.1617 | -9.3116
5.2619 | -34.162 | -9.3222
5.268 | -34.1623 | -9.3327
5.2741 | -34.1627 | -9.3431
5.2802 | -34.1629 | -9.3538
5.2863 | -34.1633 | -9.3639
5.2924 | -34.1636 | -9.3744
5.2985 | -34.1639 | -9.3848
5.3046 | -34.1642 | -9.3952
5.3107 | -34.1645 | -9.4058
5.3168 | -34.1648 | -9.4162
5.3229 | -34.165 | -9.4264
5.329 | -34.1653 | -9.4367
5.3351 | -34.1657 | -9.4465
5.3412 | -34.1659 | -9.4567
5.3473 | -34.1662 | -9.4671
5.3534 | -34.1665 | -9.4776
5.3595 | -34.1668 | -9.4879
5.3656 | -34.1672 | -9.4982
5.3717 | -34.1674 | -9.509
5.3779 | -34.1677 | -9.5199
5.384 | -34.168 | -9.5301
5.3901 | -34.1684 | -9.5409
5.3962 | -34.1686 | -9.5509
5.4023 | -34.169 | -9.5616
5.4084 | -34.1693 | -9.5718
5.4145 | -34.1697 | -9.5818
5.4206 | -34.1699 | -9.5922
5.4267 | -34.1702 | -9.6031
5.4328 | -34.1705 | -9.6135
5.4389 | -34.1708 | -9.6234
5.445 | -34.1711 | -9.6343
5.4511 | -34.1714 | -9.6444
5.4572 | -34.1718 | -9.6549
5.4633 | -34.172 | -9.6653
5.4694 | -34.1723 | -9.6757
5.4755 | -34.1726 | -9.6856
5.4816 | -34.1729 | -9.6961
5.4877 | -34.1733 | -9.7063
5.4938 | -34.1735 | -9.7163
5.4999 | -34.1739 | -9.7265
5.506 | -34.1743 | -9.7369
5.5121 | -34.1745 | -9.747
5.5183 | -34.1748 | -9.7571
5.5244 | -34.1751 | -9.7678
5.5305 | -34.1754 | -9.7781
5.5366 | -34.1757 | -9.7881
5.5427 | -34.176 | -9.7987
5.5488 | -34.1763 | -9.809
5.5549 | -34.1765 | -9.8195
5.561 | -34.1769 | -9.8302
5.5671 | -34.1772 | -9.8403
5.5732 | -34.1775 | -9.851
5.5793 | -34.1778 | -9.862
5.5854 | -34.1782 | -9.8723
5.5915 | -34.1785 | -9.8832
5.5976 | -34.1788 | -9.8937
5.6037 | -34.179 | -9.9043
5.6098 | -34.1793 | -9.9146
5.6159 | -34.1797 | -9.9248
5.622 | -34.1799 | -9.9353
5.6281 | -34.1803 | -9.9459
5.6342 | -34.1806 | -9.9569
5.6403 | -34.1809 | -9.9673
5.6464 | -34.1812 | -9.978
5.6525 | -34.1815 | -9.9885
5.6586 | -34.1818 | -9.9985
5.6648 | -34.1822 | -10.009
5.6709 | -34.1825 | -10.0197
5.677 | -34.1828 | -10.0302
5.6831 | -34.1831 | -10.0408
5.6892 | -34.1834 | -10.0515
5.6953 | -34.1837 | -10.062
5.7014 | -34.1841 | -10.0727
5.7075 | -34.1843 | -10.083
5.7136 | -34.1846 | -10.0934
5.7197 | -34.1849 | -10.1038
5.7258 | -34.1852 | -10.114
5.7319 | -34.1855 | -10.1246
5.738 | -34.1858 | -10.135
5.7441 | -34.1862 | -10.1454
5.7502 | -34.1865 | -10.1554
5.7563 | -34.1869 | -10.1659
5.7624 | -34.1872 | -10.1767
5.7685 | -34.1875 | -10.1873
5.7746 | -34.1878 | -10.1981
5.7807 | -34.188 | -10.2086
5.7868 | -34.1883 | -10.2194
5.7929 | -34.1886 | -10.2297
5.799 | -34.1889 | -10.24
5.8052 | -34.1892 | -10.2502
5.8113 | -34.1895 | -10.2605
5.8174 | -34.1898 | -10.2707
5.8235 | -34.1901 | -10.2811
5.8296 | -34.1904 | -10.2918
5.8357 | -34.1906 | -10.3021
5.8418 | -34.1909 | -10.3128
5.8479 | -34.1912 | -10.3233
5.854 | -34.1915 | -10.334
5.8601 | -34.1918 | -10.3443
5.8662 | -34.1921 | -10.3549
5.8723 | -34.1924 | -10.3652
5.8784 | -34.1927 | -10.3755
5.8845 | -34.193 | -10.3857
5.8906 | -34.1933 | -10.3963
5.8967 | -34.1936 | -10.4064
5.9028 | -34.1939 | -10.4167
5.9089 | -34.1943 | -10.427
5.915 | -34.1946 | -10.437
5.9211 | -34.1949 | -10.4474
5.9272 | -34.1952 | -10.4575
5.9333 | -34.1955 | -10.4677
5.9394 | -34.1959 | -10.4781
5.9455 | -34.1962 | -10.4882
5.9517 | -34.1966 | -10.4983
5.9578 | -34.1969 | -10.5089
5.9639 | -34.1972 | -10.5188
5.97 | -34.1976 | -10.5293
5.9761 | -34.198 | -10.5396
5.9822 | -34.1982 | -10.5501
5.9883 | -34.1986 | -10.5604
5.9944 | -34.1989 | -10.571
6.0005 | -34.1993 | -10.5816
6.0066 | -34.1996 | -10.5917
6.0127 | -34.1999 | -10.6017
6.0188 | -34.2002 | -10.6124
6.0249 | -34.2005 | -10.6228
6.031 | -34.2008 | -10.6329
6.0371 | -34.2011 | -10.6432
6.0432 | -34.2014 | -10.6534
6.0493 | -34.2018 | -10.6634
6.0554 | -34.2022 | -10.6735
6.0615 | -34.2025 | -10.6838
6.0676 | -34.2028 | -10.6945
6.0737 | -34.2032 | -10.705
6.0798 | -34.2034 | -10.7153
6.0859 | -34.2038 | -10.7256
6.0921 | -34.2041 | -10.7355
6.0982 | -34.2045 | -10.7458
6.1043 | -34.2048 | -10.756
6.1104 | -34.2052 | -10.7666
6.1165 | -34.2055 | -10.7762
6.1226 | -34.2059 | -10.7864
6.1287 | -34.2062 | -10.7965
6.1348 | -34.2065 | -10.8068
6.1409 | -34.2069 | -10.8169
6.147 | -34.2072 | -10.8271
6.1531 | -34.2075 | -10.8372
6.1592 | -34.2078 | -10.8473
6.1653 | -34.2081 | -10.8574
6.1714 | -34.2084 | -10.8676
6.1775 | -34.2088 | -10.878
6.1836 | -34.2091 | -10.8882
6.1897 | -34.2094 | -10.8985
6.1958 | -34.2097 | -10.9091
6.2019 | -34.2101 | -10.9195
6.208 | -34.2105 | -10.9298
6.2141 | -34.2108 | -10.9399
6.2202 | -34.2112 | -10.9504
6.2263 | -34.2115 | -10.9605
6.2325 | -34.2118 | -10.9702
6.2386 | -34.2121 | -10.9799
6.2447 | -34.2125 | -10.9905
6.2508 | -34.2128 | -11.0007
6.2569 | -34.2132 | -11.0108
6.263 | -34.2135 | -11.021
6.2691 | -34.2138 | -11.031
6.2752 | -34.2141 | -11.0413
6.2813 | -34.2144 | -11.0515
6.2874 | -34.2148 | -11.0615
6.2935 | -34.2151 | -11.0716
6.2996 | -34.2154 | -11.0816
6.3057 | -34.2157 | -11.0919
6.3118 | -34.216 | -11.1022
6.3179 | -34.2163 | -11.1123
6.324 | -34.2167 | -11.1224
6.3301 | -34.217 | -11.1325
6.3362 | -34.2174 | -11.1425
6.3423 | -34.2177 | -11.153
6.3484 | -34.2181 | -11.1632
6.3545 | -34.2183 | -11.1729
6.3606 | -34.2187 | -11.1826
6.3667 | -34.2189 | -11.1925
6.3728 | -34.2193 | -11.2026
6.379 | -34.2196 | -11.2128
6.3851 | -34.22 | -11.2225
6.3912 | -34.2202 | -11.2325
6.3973 | -34.2205 | -11.2426
6.4034 | -34.2208 | -11.2525
6.4095 | -34.2212 | -11.2626
6.4156 | -34.2215 | -11.2727
6.4217 | -34.2218 | -11.2827
6.4278 | -34.2221 | -11.2929
6.4339 | -34.2225 | -11.3028
6.44 | -34.2229 | -11.3127
6.4461 | -34.2231 | -11.3231
6.4522 | -34.2235 | -11.3328
6.4583 | -34.2237 | -11.343
6.4644 | -34.224 | -11.3533
6.4705 | -34.2243 | -11.3628
6.4766 | -34.2247 | -11.373
6.4827 | -34.225 | -11.3831
6.4888 | -34.2252 | -11.3928
6.4949 | -34.2256 | -11.4024
6.501 | -34.2259 | -11.4126
6.5071 | -34.2262 | -11.4223
6.5132 | -34.2266 | -11.4324
6.5194 | -34.2268 | -11.4425
6.5255 | -34.2272 | -11.4521
6.5316 | -34.2274 | -11.4623
6.5377 | -34.2278 | -11.4723
6.5438 | -34.228 | -11.482
6.5499 | -34.2284 | -11.4918
6.556 | -34.2287 | -11.5018
6.5621 | -34.229 | -11.5121
6.5682 | -34.2293 | -11.5222
6.5743 | -34.2297 | -11.5321
6.5804 | -34.23 | -11.5423
6.5865 | -34.2303 | -11.5523
6.5926 | -34.2306 | -11.5621
6.5987 | -34.2309 | -11.5717
6.6048 | -34.2313 | -11.5817
6.6109 | -34.2316 | -11.5917
6.617 | -34.2318 | -11.6018
6.6231 | -34.2321 | -11.6117
6.6292 | -34.2325 | -11.6217
6.6353 | -34.2328 | -11.6321
6.6414 | -34.2331 | -11.6419
6.6475 | -34.2335 | -11.6518
6.6536 | -34.2337 | -11.6616
6.6597 | -34.234 | -11.6711
6.6659 | -34.2344 | -11.6811
6.672 | -34.2347 | -11.6908
6.6781 | -34.235 | -11.7008
6.6842 | -34.2353 | -11.7105
6.6903 | -34.2357 | -11.7204
6.6964 | -34.2361 | -11.7304
6.7025 | -34.2363 | -11.74
6.7086 | -34.2366 | -11.7498
6.7147 | -34.2369 | -11.7599
6.7208 | -34.2372 | -11.77
6.7269 | -34.2375 | -11.7798
6.733 | -34.2378 | -11.7894
6.7391 | -34.2382 | -11.7993
6.7452 | -34.2385 | -11.8088
6.7513 | -34.2388 | -11.8184
6.7574 | -34.2391 | -11.8284
6.7635 | -34.2394 | -11.8379
6.7696 | -34.2398 | -11.8476
6.7757 | -34.2402 | -11.8575
6.7818 | -34.2405 | -11.8674
6.7879 | -34.2408 | -11.8771
6.794 | -34.2411 | -11.8866
6.8001 | -34.2414 | -11.8963
6.8063 | -34.2417 | -11.9061
6.8124 | -34.242 | -11.9162
6.8185 | -34.2423 | -11.9263
6.8246 | -34.2426 | -11.9359
6.8307 | -34.2429 | -11.9455
6.8368 | -34.2432 | -11.9557
6.8429 | -34.2436 | -11.9656
6.849 | -34.2439 | -11.9754
6.8551 | -34.2443 | -11.9851
6.8612 | -34.2445 | -11.9949
6.8673 | -34.2448 | -12.0046
6.8734 | -34.2451 | -12.0147
6.8795 | -34.2455 | -12.0246
6.8856 | -34.2458 | -12.0344
6.8917 | -34.2461 | -12.0441
6.8978 | -34.2464 | -12.0535
6.9039 | -34.2467 | -12.0629
6.91 | -34.2471 | -12.0729
6.9161 | -34.2474 | -12.0823
6.9222 | -34.2477 | -12.0921
6.9283 | -34.2481 | -12.1015
6.9344 | -34.2484 | -12.1113
6.9405 | -34.2488 | -12.1211
6.9466 | -34.2491 | -12.1312
6.9528 | -34.2495 | -12.1412
6.9589 | -34.2498 | -12.1512
6.965 | -34.2502 | -12.1611
6.9711 | -34.2505 | -12.1711
6.9772 | -34.2509 | -12.1809
6.9833 | -34.2512 | -12.1906
6.9894 | -34.2515 | -12.2007
6.9955 | -34.2518 | -12.2102
7.0016 | -34.2522 | -12.2199
7.0077 | -34.2525 | -12.2296
7.0138 | -34.2529 | -12.2393
7.0199 | -34.2531 | -12.2489
7.026 | -34.2535 | -12.2586
7.0321 | -34.2538 | -12.268
7.0382 | -34.2542 | -12.2773
7.0443 | -34.2545 | -12.2873
7.0504 | -34.2548 | -12.2971
7.0565 | -34.2552 | -12.3069
7.0626 | -34.2555 | -12.3162
7.0687 | -34.2558 | -12.3254
7.0748 | -34.2562 | -12.3347
7.0809 | -34.2565 | -12.3444
7.087 | -34.2569 | -12.3542
7.0932 | -34.2572 | -12.364
7.0993 | -34.2575 | -12.3737
7.1054 | -34.2578 | -12.3833
7.1115 | -34.2582 | -12.3926
7.1176 | -34.2585 | -12.4022
7.1237 | -34.2588 | -12.4117
7.1298 | -34.2592 | -12.4213
7.1359 | -34.2595 | -12.4309
7.142 | -34.2598 | -12.4402
7.1481 | -34.2601 | -12.4497
7.1542 | -34.2604 | -12.4594
7.1603 | -34.2608 | -12.4691
7.1664 | -34.2612 | -12.4787
7.1725 | -34.2615 | -12.4884
7.1786 | -34.2619 | -12.4981
7.1847 | -34.2622 | -12.5078
7.1908 | -34.2625 | -12.5176
7.1969 | -34.2628 | -12.5275
7.203 | -34.2631 | -12.5369
7.2091 | -34.2635 | -12.5464
7.2152 | -34.2638 | -12.5561
7.2213 | -34.2642 | -12.5654
7.2274 | -34.2645 | -12.5749
7.2335 | -34.2648 | -12.5845
7.2397 | -34.2652 | -12.594
7.2458 | -34.2656 | -12.6033
7.2519 | -34.266 | -12.613
7.258 | -34.2663 | -12.6229
7.2641 | -34.2666 | -12.6325
7.2702 | -34.2669 | -12.6424
7.2763 | -34.2673 | -12.652
7.2824 | -34.2677 | -12.6619
7.2885 | -34.268 | -12.6716
7.2946 | -34.2683 | -12.6812
7.3007 | -34.2686 | -12.691
7.3068 | -34.2689 | -12.7009
7.3129 | -34.2693 | -12.7109
7.319 | -34.2697 | -12.7205
7.3251 | -34.27 | -12.7299
7.3312 | -34.2703 | -12.7395
7.3373 | -34.2707 | -12.7488
7.3434 | -34.2711 | -12.7583
7.3495 | -34.2715 | -12.7678
7.3556 | -34.2718 | -12.7775
7.3617 | -34.2722 | -12.7875
7.3678 | -34.2725 | -12.7969
7.3739 | -34.2728 | -12.8068
7.3801 | -34.2731 | -12.8169
7.3862 | -34.2735 | -12.8265
7.3923 | -34.2739 | -12.8363
7.3984 | -34.2742 | -12.8456
7.4045 | -34.2746 | -12.8552
7.4106 | -34.275 | -12.8648
7.4167 | -34.2753 | -12.8745
7.4228 | -34.2757 | -12.8844
7.4289 | -34.276 | -12.894
7.435 | -34.2764 | -12.9033
7.4411 | -34.2767 | -12.9131
7.4472 | -34.2771 | -12.9228
7.4533 | -34.2774 | -12.9322
7.4594 | -34.2778 | -12.9422
7.4655 | -34.2782 | -12.9516
7.4716 | -34.2785 | -12.9615
7.4777 | -34.2788 | -12.9711
7.4838 | -34.2792 | -12.9805
7.4899 | -34.2795 | -12.99
7.496 | -34.2799 | -12.9994
7.5021 | -34.2802 | -13.009
7.5082 | -34.2806 | -13.0187
7.5143 | -34.2809 | -13.0284
7.5204 | -34.2813 | -13.0382
7.5266 | -34.2816 | -13.0478
7.5327 | -34.282 | -13.0573
7.5388 | -34.2823 | -13.0674
7.5449 | -34.2827 | -13.0773
7.551 | -34.283 | -13.0868
7.5571 | -34.2834 | -13.0965
7.5632 | -34.2838 | -13.1063
7.5693 | -34.2842 | -13.1158
7.5754 | -34.2845 | -13.1252
7.5815 | -34.2848 | -13.1349
7.5876 | -34.2852 | -13.1449
7.5937 | -34.2856 | -13.1549
7.5998 | -34.286 | -13.1645
7.6059 | -34.2864 | -13.1745
7.612 | -34.2867 | -13.1841
7.6181 | -34.2871 | -13.1936
7.6242 | -34.2874 | -13.2036
7.6303 | -34.2878 | -13.2137
7.6364 | -34.2881 | -13.2235
7.6425 | -34.2884 | -13.2333
7.6486 | -34.2888 | -13.243
7.6547 | -34.2891 | -13.2526
7.6608 | -34.2895 | -13.2623
7.667 | -34.2898 | -13.2722
7.6731 | -34.2902 | -13.282
7.6792 | -34.2906 | -13.292
7.6853 | -34.2909 | -13.3019
7.6914 | -34.2913 | -13.3118
7.6975 | -34.2916 | -13.3216
7.7036 | -34.292 | -13.3315
7.7097 | -34.2924 | -13.3415
7.7158 | -34.2928 | -13.3513
7.7219 | -34.2931 | -13.3608
7.728 | -34.2935 | -13.3708
7.7341 | -34.2938 | -13.3807
7.7402 | -34.2942 | -13.3905
7.7463 | -34.2946 | -13.4004
7.7524 | -34.2949 | -13.4105
7.7585 | -34.2953 | -13.4204
7.7646 | -34.2956 | -13.4301
7.7707 | -34.296 | -13.4402
7.7768 | -34.2963 | -13.4502
7.7829 | -34.2966 | -13.4604
7.789 | -34.297 | -13.4706
7.7951 | -34.2973 | -13.4803
7.8012 | -34.2977 | -13.4902
7.8073 | -34.2981 | -13.5004
7.8135 | -34.2984 | -13.5102
7.8196 | -34.2989 | -13.5204
7.8257 | -34.2992 | -13.5305
7.8318 | -34.2995 | -13.5405
7.8379 | -34.2999 | -13.5502
7.844 | -34.3002 | -13.56
7.8501 | -34.3006 | -13.5701
7.8562 | -34.3009 | -13.5803
7.8623 | -34.3014 | -13.5901
7.8684 | -34.3017 | -13.6
7.8745 | -34.302 | -13.6099
7.8806 | -34.3024 | -13.6201
7.8867 | -34.3028 | -13.6297
7.8928 | -34.3031 | -13.6393
7.8989 | -34.3035 | -13.6493
7.905 | -34.3038 | -13.6593
7.9111 | -34.3042 | -13.6693
7.9172 | -34.3045 | -13.6792
7.9233 | -34.3048 | -13.6892
7.9294 | -34.3052 | -13.6992
7.9355 | -34.3056 | -13.7088
7.9416 | -34.3059 | -13.7188
7.9477 | -34.3063 | -13.7285
7.9539 | -34.3066 | -13.7382
7.96 | -34.307 | -13.7483
7.9661 | -34.3073 | -13.7583
7.9722 | -34.3078 | -13.7679
7.9783 | -34.3081 | -13.7782
7.9844 | -34.3085 | -13.7881
7.9905 | -34.3088 | -13.7983
7.9966 | -34.3092 | -13.8084
8.0027 | -34.3095 | -13.8183
8.0088 | -34.3099 | -13.8282
8.0149 | -34.3102 | -13.8381
8.021 | -34.3105 | -13.8482
8.0271 | -34.3109 | -13.8586
8.0332 | -34.3113 | -13.8683
8.0393 | -34.3117 | -13.8785
8.0454 | -34.312 | -13.8889
8.0515 | -34.3124 | -13.8989
8.0576 | -34.3128 | -13.9092
8.0637 | -34.3131 | -13.919
8.0698 | -34.3134 | -13.9288
8.0759 | -34.3137 | -13.9386
8.082 | -34.3141 | -13.9484
8.0881 | -34.3144 | -13.9585
8.0942 | -34.3148 | -13.9686
8.1004 | -34.3153 | -13.9787
8.1065 | -34.3156 | -13.9884
8.1126 | -34.3159 | -13.9984
8.1187 | -34.3162 | -14.0086
8.1248 | -34.3166 | -14.0182
8.1309 | -34.317 | -14.028
8.137 | -34.3173 | -14.0377
8.1431 | -34.3177 | -14.048
8.1492 | -34.318 | -14.0579
8.1553 | -34.3184 | -14.068
8.1614 | -34.3187 | -14.0782
8.1675 | -34.3191 | -14.0882
8.1736 | -34.3194 | -14.0984
8.1797 | -34.3197 | -14.1085
8.1858 | -34.3201 | -14.1186
8.1919 | -34.3204 | -14.1286
8.198 | -34.3208 | -14.1383
8.2041 | -34.3212 | -14.1484
8.2102 | -34.3215 | -14.1583
8.2163 | -34.3219 | -14.168
8.2224 | -34.3222 | -14.178
8.2285 | -34.3226 | -14.1876
8.2346 | -34.3229 | -14.1975
8.2408 | -34.3233 | -14.2073
8.2469 | -34.3237 | -14.2174
8.253 | -34.3241 | -14.2271
8.2591 | -34.3244 | -14.2372
8.2652 | -34.3248 | -14.2472
8.2713 | -34.3252 | -14.2573
8.2774 | -34.3255 | -14.2675
8.2835 | -34.3259 | -14.2778
8.2896 | -34.3262 | -14.2875
8.2957 | -34.3266 | -14.2976
8.3018 | -34.3269 | -14.3076
8.3079 | -34.3273 | -14.3174
8.314 | -34.3277 | -14.3274
8.3201 | -34.3281 | -14.3373
8.3262 | -34.3284 | -14.3473
8.3323 | -34.3287 | -14.3572
8.3384 | -34.3291 | -14.3674
8.3445 | -34.3294 | -14.3773
8.3506 | -34.3298 | -14.3873
8.3567 | -34.3302 | -14.3974
8.3628 | -34.3306 | -14.4073
8.3689 | -34.3309 | -14.4171
8.375 | -34.3313 | -14.4273
8.3812 | -34.3316 | -14.437
8.3873 | -34.332 | -14.4469
8.3934 | -34.3323 | -14.4571
8.3995 | -34.3326 | -14.4668
8.4056 | -34.333 | -14.4767
8.4117 | -34.3333 | -14.4868
8.4178 | -34.3337 | -14.4968
8.4239 | -34.334 | -14.5065
8.43 | -34.3344 | -14.5165
8.4361 | -34.3347 | -14.5265
8.4422 | -34.3351 | -14.5363
8.4483 | -34.3354 | -14.5462
8.4544 | -34.3358 | -14.5561
8.4605 | -34.3362 | -14.5659
8.4666 | -34.3365 | -14.5762
8.4727 | -34.3368 | -14.5861
8.4788 | -34.3372 | -14.596
8.4849 | -34.3375 | -14.6059
8.491 | -34.3379 | -14.6159
8.4971 | -34.3382 | -14.6255
8.5032 | -34.3386 | -14.6352
8.5093 | -34.3389 | -14.645
8.5154 | -34.3393 | -14.655
8.5215 | -34.3397 | -14.6648
8.5277 | -34.34 | -14.6747
8.5338 | -34.3404 | -14.6845
8.5399 | -34.3407 | -14.6944
8.546 | -34.3411 | -14.7042
8.5521 | -34.3415 | -14.7141
8.5582 | -34.3418 | -14.7238
8.5643 | -34.3422 | -14.7333
8.5704 | -34.3425 | -14.7429
8.5765 | -34.3429 | -14.7529
8.5826 | -34.3432 | -14.7623
8.5887 | -34.3436 | -14.7721
8.5948 | -34.344 | -14.782
8.6009 | -34.3444 | -14.7916
8.607 | -34.3447 | -14.8017
8.6131 | -34.3451 | -14.8116
8.6192 | -34.3455 | -14.8213
8.6253 | -34.3459 | -14.831
8.6314 | -34.3462 | -14.8408
8.6375 | -34.3465 | -14.8506
8.6436 | -34.3469 | -14.8605
8.6497 | -34.3472 | -14.8701
8.6558 | -34.3475 | -14.8799
8.6619 | -34.3479 | -14.8897
8.6681 | -34.3482 | -14.8998
8.6742 | -34.3486 | -14.9095
8.6803 | -34.349 | -14.9193
8.6864 | -34.3494 | -14.9287
8.6925 | -34.3497 | -14.9382
8.6986 | -34.3501 | -14.9482
8.7047 | -34.3504 | -14.9583
8.7108 | -34.3508 | -14.9681
8.7169 | -34.3511 | -14.9778
8.723 | -34.3515 | -14.9873
8.7291 | -34.3518 | -14.9969
8.7352 | -34.3522 | -15.007
8.7413 | -34.3525 | -15.0168
8.7474 | -34.3529 | -15.0263
8.7535 | -34.3533 | -15.0361
8.7596 | -34.3536 | -15.0461
8.7657 | -34.354 | -15.056
8.7718 | -34.3543 | -15.0659
8.7779 | -34.3547 | -15.0759
8.784 | -34.355 | -15.0858
8.7901 | -34.3555 | -15.0956
8.7962 | -34.3558 | -15.105
8.8023 | -34.3562 | -15.1147
8.8084 | -34.3565 | -15.1243
8.8146 | -34.3569 | -15.134
8.8207 | -34.3572 | -15.1436
8.8268 | -34.3576 | -15.153
8.8329 | -34.358 | -15.1626
8.839 | -34.3584 | -15.1724
8.8451 | -34.3587 | -15.1819
8.8512 | -34.3591 | -15.1915
8.8573 | -34.3594 | -15.2012
8.8634 | -34.3598 | -15.2107
8.8695 | -34.3602 | -15.2207
8.8756 | -34.3605 | -15.2305
8.8817 | -34.3609 | -15.2401
8.8878 | -34.3612 | -15.2501
8.8939 | -34.3616 | -15.2598
8.9 | -34.362 | -15.2695
8.9061 | -34.3623 | -15.2791
8.9122 | -34.3627 | -15.2889
8.9183 | -34.3631 | -15.2987
8.9244 | -34.3635 | -15.3086
8.9305 | -34.3639 | -15.318
8.9366 | -34.3643 | -15.3278
8.9427 | -34.3646 | -15.3375
8.9488 | -34.365 | -15.3471
8.955 | -34.3653 | -15.3566
8.9611 | -34.3657 | -15.3666
8.9672 | -34.3662 | -15.3762
8.9733 | -34.3665 | -15.3857
8.9794 | -34.367 | -15.3952
8.9855 | -34.3673 | -15.405
8.9916 | -34.3677 | -15.4149
8.9977 | -34.368 | -15.4245
9.0038 | -34.3685 | -15.4342
9.0099 | -34.3689 | -15.4439
9.016 | -34.3692 | -15.4536
9.0221 | -34.3696 | -15.463
9.0282 | -34.37 | -15.4724
9.0343 | -34.3704 | -15.4821
9.0404 | -34.3708 | -15.4918
9.0465 | -34.3712 | -15.5018
9.0526 | -34.3716 | -15.5113
9.0587 | -34.372 | -15.5211
9.0648 | -34.3724 | -15.5306
9.0709 | -34.3728 | -15.5401
9.077 | -34.3732 | -15.5499
9.0831 | -34.3736 | -15.5598
9.0892 | -34.374 | -15.5696
9.0953 | -34.3744 | -15.5794
9.1015 | -34.3748 | -15.589
9.1076 | -34.3752 | -15.5984
9.1137 | -34.3756 | -15.608
9.1198 | -34.376 | -15.6176
9.1259 | -34.3764 | -15.6273
9.132 | -34.3768 | -15.6372
9.1381 | -34.3773 | -15.6467
9.1442 | -34.3777 | -15.6559
9.1503 | -34.3781 | -15.6658
9.1564 | -34.3785 | -15.6753
9.1625 | -34.3788 | -15.6848
9.1686 | -34.3793 | -15.6944
9.1747 | -34.3797 | -15.704
9.1808 | -34.3801 | -15.7137
9.1869 | -34.3805 | -15.7233
9.193 | -34.3809 | -15.7329
9.1991 | -34.3813 | -15.7427
9.2052 | -34.3817 | -15.7522
9.2113 | -34.3821 | -15.7618
9.2174 | -34.3826 | -15.7714
9.2235 | -34.383 | -15.781
9.2296 | -34.3834 | -15.7906
9.2357 | -34.3838 | -15.8005
9.2419 | -34.3842 | -15.8101
9.248 | -34.3847 | -15.8198
9.2541 | -34.3851 | -15.8295
9.2602 | -34.3856 | -15.8393
9.2663 | -34.386 | -15.8486
9.2724 | -34.3864 | -15.8581
9.2785 | -34.3868 | -15.8676
9.2846 | -34.3872 | -15.877
9.2907 | -34.3876 | -15.8865
9.2968 | -34.388 | -15.8964
9.3029 | -34.3884 | -15.9062
9.309 | -34.3888 | -15.9159
9.3151 | -34.3892 | -15.9252
9.3212 | -34.3897 | -15.9347
9.3273 | -34.3901 | -15.9442
9.3334 | -34.3905 | -15.9536
9.3395 | -34.391 | -15.9629
9.3456 | -34.3913 | -15.9722
9.3517 | -34.3917 | -15.9819
9.3578 | -34.3922 | -15.9915
9.3639 | -34.3926 | -16.0008
9.37 | -34.393 | -16.0105
9.3761 | -34.3934 | -16.0203
9.3822 | -34.3939 | -16.03
9.3884 | -34.3943 | -16.0391
9.3945 | -34.3947 | -16.0488
9.4006 | -34.3952 | -16.0584
9.4067 | -34.3956 | -16.0681
9.4128 | -34.396 | -16.0778
9.4189 | -34.3965 | -16.0877
9.425 | -34.3969 | -16.0971
9.4311 | -34.3974 | -16.1068
9.4372 | -34.3978 | -16.1163
9.4433 | -34.3982 | -16.1259
9.4494 | -34.3987 | -16.1352
9.4555 | -34.3991 | -16.145
9.4616 | -34.3996 | -16.1545
9.4677 | -34.4 | -16.1641
9.4738 | -34.4004 | -16.1737
9.4799 | -34.4008 | -16.1833
9.486 | -34.4012 | -16.1929
9.4921 | -34.4017 | -16.2024
9.4982 | -34.4021 | -16.2117
9.5043 | -34.4026 | -16.2211
9.5104 | -34.403 | -16.2304
9.5165 | -34.4034 | -16.24
9.5226 | -34.4039 | -16.2497
9.5288 | -34.4043 | -16.2591
9.5349 | -34.4047 | -16.2688
9.541 | -34.4052 | -16.278
9.5471 | -34.4056 | -16.2875
9.5532 | -34.406 | -16.2972
9.5593 | -34.4065 | -16.307
9.5654 | -34.4069 | -16.3165
9.5715 | -34.4073 | -16.3259
9.5776 | -34.4077 | -16.3357
9.5837 | -34.4081 | -16.3451
9.5898 | -34.4085 | -16.3546
9.5959 | -34.409 | -16.3643
9.602 | -34.4094 | -16.3736
9.6081 | -34.4098 | -16.3831
9.6142 | -34.4103 | -16.3925
9.6203 | -34.4107 | -16.4021
9.6264 | -34.4111 | -16.4115
9.6325 | -34.4115 | -16.4212
9.6386 | -34.4119 | -16.4306
9.6447 | -34.4124 | -16.4401
9.6508 | -34.4128 | -16.4495
9.6569 | -34.4132 | -16.459
9.663 | -34.4137 | -16.4687
9.6691 | -34.4141 | -16.4782
9.6753 | -34.4146 | -16.4877
9.6814 | -34.415 | -16.4971
9.6875 | -34.4154 | -16.5063
9.6936 | -34.4159 | -16.5155
9.6997 | -34.4163 | -16.5252
9.7058 | -34.4168 | -16.5346
9.7119 | -34.4172 | -16.5444
9.718 | -34.4177 | -16.5541
9.7241 | -34.4181 | -16.5634
9.7302 | -34.4185 | -16.573
9.7363 | -34.419 | -16.5823
9.7424 | -34.4194 | -16.5916
9.7485 | -34.4198 | -16.6012
9.7546 | -34.4202 | -16.6104
9.7607 | -34.4207 | -16.6202
9.7668 | -34.4211 | -16.6298
9.7729 | -34.4216 | -16.6392
9.779 | -34.422 | -16.6485
9.7851 | -34.4225 | -16.6582
9.7912 | -34.423 | -16.6676
9.7973 | -34.4234 | -16.6773
9.8034 | -34.4239 | -16.6865
9.8095 | -34.4243 | -16.6963
9.8157 | -34.4248 | -16.7058
9.8218 | -34.4252 | -16.7156
9.8279 | -34.4256 | -16.7248
9.834 | -34.426 | -16.7344
9.8401 | -34.4265 | -16.7439
9.8462 | -34.427 | -16.7531
9.8523 | -34.4274 | -16.7625
9.8584 | -34.4278 | -16.772
9.8645 | -34.4282 | -16.7815
9.8706 | -34.4286 | -16.7911
9.8767 | -34.429 | -16.8004
9.8828 | -34.4294 | -16.8097
9.8889 | -34.4299 | -16.8189
9.895 | -34.4303 | -16.8284
9.9011 | -34.4307 | -16.8379
9.9072 | -34.4312 | -16.8473
9.9133 | -34.4316 | -16.8567
9.9194 | -34.432 | -16.8661
9.9255 | -34.4324 | -16.8755
9.9316 | -34.4329 | -16.8849
9.9377 | -34.4333 | -16.8943
9.9438 | -34.4337 | -16.9038
9.9499 | -34.4342 | -16.9135
9.956 | -34.4346 | -16.9229
9.9622 | -34.435 | -16.9325
9.9683 | -34.4355 | -16.9419
9.9744 | -34.4359 | -16.9516
9.9805 | -34.4364 | -16.961
9.9866 | -34.4367 | -16.9703
9.9927 | -34.4372 | -16.9796
9.9988 | -34.4376 | -16.9892
10.0049 | -34.438 | -16.9986
10.011 | -34.4385 | -17.0082
10.0171 | -34.4389 | -17.0175
10.0232 | -34.4394 | -17.0268
10.0293 | -34.4398 | -17.0364
10.0354 | -34.4403 | -17.0459
10.0415 | -34.4408 | -17.0554
10.0476 | -34.4412 | -17.065
10.0537 | -34.4416 | -17.0746
10.0598 | -34.442 | -17.084
10.0659 | -34.4424 | -17.0935
10.072 | -34.4428 | -17.1026
10.0781 | -34.4433 | -17.1119
10.0842 | -34.4437 | -17.1213
10.0903 | -34.4442 | -17.1307
10.0964 | -34.4446 | -17.1398
10.1026 | -34.445 | -17.1492
10.1087 | -34.4454 | -17.1586
10.1148 | -34.4458 | -17.1681
10.1209 | -34.4463 | -17.1774
10.127 | -34.4467 | -17.1866
10.1331 | -34.4471 | -17.1958
10.1392 | -34.4475 | -17.2049
10.1453 | -34.448 | -17.2143
10.1514 | -34.4484 | -17.2238
10.1575 | -34.4488 | -17.2334
10.1636 | -34.4493 | -17.2424
10.1697 | -34.4497 | -17.2519
10.1758 | -34.4501 | -17.261
10.1819 | -34.4506 | -17.2703
10.188 | -34.4509 | -17.2797
10.1941 | -34.4514 | -17.2891
10.2002 | -34.4518 | -17.2985
10.2063 | -34.4522 | -17.3079
10.2124 | -34.4526 | -17.3174
10.2185 | -34.4531 | -17.3269
10.2246 | -34.4535 | -17.3361
10.2307 | -34.4539 | -17.3455
10.2368 | -34.4543 | -17.3544
10.2429 | -34.4547 | -17.3637
10.2491 | -34.4551 | -17.3732
10.2552 | -34.4556 | -17.3826
10.2613 | -34.456 | -17.3921
10.2674 | -34.4564 | -17.4015
10.2735 | -34.4568 | -17.4112
10.2796 | -34.4573 | -17.4205
10.2857 | -34.4577 | -17.4301
10.2918 | -34.4581 | -17.4397
10.2979 | -34.4585 | -17.449
10.304 | -34.459 | -17.4581
10.3101 | -34.4594 | -17.4674
10.3162 | -34.4598 | -17.477
10.3223 | -34.4602 | -17.4861
10.3284 | -34.4607 | -17.4953
10.3345 | -34.4611 | -17.5045
10.3406 | -34.4615 | -17.5138
10.3467 | -34.4618 | -17.5231
10.3528 | -34.4623 | -17.5322
10.3589 | -34.4627 | -17.5416
10.365 | -34.4631 | -17.5511
10.3711 | -34.4635 | -17.5605
10.3772 | -34.4639 | -17.5699
10.3833 | -34.4644 | -17.5793
10.3895 | -34.4648 | -17.5886
10.3956 | -34.4652 | -17.5979
10.4017 | -34.4657 | -17.607
10.4078 | -34.4661 | -17.6163
10.4139 | -34.4665 | -17.6257
10.42 | -34.4669 | -17.6353
10.4261 | -34.4673 | -17.6446
10.4322 | -34.4677 | -17.6539
10.4383 | -34.4681 | -17.6632
10.4444 | -34.4685 | -17.6723
10.4505 | -34.4689 | -17.6816
10.4566 | -34.4693 | -17.6907
10.4627 | -34.4697 | -17.6998
10.4688 | -34.4702 | -17.709
10.4749 | -34.4706 | -17.7183
10.481 | -34.471 | -17.7278
10.4871 | -34.4714 | -17.737
10.4932 | -34.4718 | -17.7461
10.4993 | -34.4723 | -17.7555
10.5054 | -34.4727 | -17.7648
10.5115 | -34.4731 | -17.7741
10.5176 | -34.4735 | -17.7833
10.5237 | -34.474 | -17.7929
10.5298 | -34.4744 | -17.8019
10.536 | -34.4747 | -17.8113
10.5421 | -34.4752 | -17.8204
10.5482 | -34.4756 | -17.8298
10.5543 | -34.476 | -17.839
10.5604 | -34.4764 | -17.848
10.5665 | -34.4768 | -17.8571
10.5726 | -34.4772 | -17.8664
10.5787 | -34.4776 | -17.8756
10.5848 | -34.478 | -17.8844
10.5909 | -34.4785 | -17.8939
10.597 | -34.4789 | -17.9032
10.6031 | -34.4793 | -17.9128
10.6092 | -34.4796 | -17.922
10.6153 | -34.4801 | -17.9314
10.6214 | -34.4805 | -17.9405
10.6275 | -34.4809 | -17.9497
10.6336 | -34.4813 | -17.959
10.6397 | -34.4817 | -17.9682
10.6458 | -34.4821 | -17.9774
10.6519 | -34.4825 | -17.9868
10.658 | -34.4829 | -17.9961
10.6641 | -34.4833 | -18.0055
10.6702 | -34.4837 | -18.0147
10.6764 | -34.4841 | -18.0242
10.6825 | -34.4846 | -18.0333
10.6886 | -34.485 | -18.0422
10.6947 | -34.4854 | -18.0515
10.7008 | -34.4858 | -18.0608
10.7069 | -34.4862 | -18.0701
10.713 | -34.4866 | -18.0793
10.7191 | -34.487 | -18.0885
10.7252 | -34.4874 | -18.0974
10.7313 | -34.4878 | -18.1065
10.7374 | -34.4882 | -18.1158
10.7435 | -34.4886 | -18.1251
10.7496 | -34.489 | -18.1341
10.7557 | -34.4894 | -18.1435
10.7618 | -34.4898 | -18.1523
10.7679 | -34.4902 | -18.1615
10.774 | -34.4906 | -18.1708
10.7801 | -34.4911 | -18.1802
10.7862 | -34.4915 | -18.1891
10.7923 | -34.4919 | -18.1984
10.7984 | -34.4923 | -18.2075
10.8045 | -34.4927 | -18.2164
10.8106 | -34.4931 | -18.2257
10.8168 | -34.4935 | -18.2348
10.8229 | -34.4939 | -18.2441
10.829 | -34.4943 | -18.2532
10.8351 | -34.4947 | -18.2621
10.8412 | -34.4952 | -18.2711
10.8473 | -34.4956 | -18.28
10.8534 | -34.496 | -18.2892
10.8595 | -34.4964 | -18.2983
10.8656 | -34.4969 | -18.3071
10.8717 | -34.4973 | -18.3163
10.8778 | -34.4977 | -18.3255
10.8839 | -34.4981 | -18.3346
10.89 | -34.4985 | -18.3438
10.8961 | -34.499 | -18.353
10.9022 | -34.4993 | -18.362
10.9083 | -34.4998 | -18.3709
10.9144 | -34.5001 | -18.38
10.9205 | -34.5005 | -18.3892
10.9266 | -34.501 | -18.3982
10.9327 | -34.5014 | -18.4072
10.9388 | -34.5018 | -18.4161
10.9449 | -34.5022 | -18.4252
10.951 | -34.5027 | -18.4343
10.9571 | -34.5031 | -18.4434
10.9633 | -34.5034 | -18.4527
10.9694 | -34.5038 | -18.4621
10.9755 | -34.5043 | -18.4713
10.9816 | -34.5047 | -18.4805
10.9877 | -34.5051 | -18.4894
10.9938 | -34.5055 | -18.4987
10.9999 | -34.506 | -18.5077
11.006 | -34.5064 | -18.5165
11.0121 | -34.5068 | -18.5256
11.0182 | -34.5072 | -18.5346
11.0243 | -34.5076 | -18.5437
11.0304 | -34.508 | -18.5528
11.0365 | -34.5084 | -18.5617
11.0426 | -34.5088 | -18.5708
11.0487 | -34.5093 | -18.5801
11.0548 | -34.5097 | -18.5889
11.0609 | -34.5101 | -18.5981
11.067 | -34.5105 | -18.6072
11.0731 | -34.511 | -18.6161
11.0792 | -34.5114 | -18.6252
11.0853 | -34.5118 | -18.6342
11.0914 | -34.5122 | -18.6435
11.0975 | -34.5127 | -18.6525
11.1037 | -34.513 | -18.6614
11.1098 | -34.5134 | -18.6707
11.1159 | -34.5139 | -18.6794
11.122 | -34.5142 | -18.6883
11.1281 | -34.5147 | -18.6974
11.1342 | -34.5151 | -18.7065
11.1403 | -34.5155 | -18.7158
11.1464 | -34.5159 | -18.7249
11.1525 | -34.5163 | -18.7336
11.1586 | -34.5167 | -18.7429
11.1647 | -34.5172 | -18.7518
11.1708 | -34.5175 | -18.7608
11.1769 | -34.5179 | -18.7699
11.183 | -34.5184 | -18.7788
11.1891 | -34.5188 | -18.7878
11.1952 | -34.5193 | -18.7968
11.2013 | -34.5197 | -18.8059
11.2074 | -34.5201 | -18.8148
11.2135 | -34.5205 | -18.8241
11.2196 | -34.5209 | -18.8331
11.2257 | -34.5214 | -18.8422
11.2318 | -34.5218 | -18.8511
11.2379 | -34.5223 | -18.8603
11.244 | -34.5227 | -18.8696
11.2502 | -34.5232 | -18.8786
11.2563 | -34.5236 | -18.8877
11.2624 | -34.524 | -18.8965
11.2685 | -34.5244 | -18.9055
11.2746 | -34.5248 | -18.9144
11.2807 | -34.5253 | -18.9236
11.2868 | -34.5257 | -18.9328
11.2929 | -34.5261 | -18.9414
11.299 | -34.5266 | -18.9505
11.3051 | -34.527 | -18.9597
11.3112 | -34.5274 | -18.9688
11.3173 | -34.5279 | -18.9776
11.3234 | -34.5283 | -18.9863
11.3295 | -34.5287 | -18.9955
11.3356 | -34.5291 | -19.0047
11.3417 | -34.5296 | -19.0135
11.3478 | -34.53 | -19.0227
11.3539 | -34.5303 | -19.0318
11.36 | -34.5308 | -19.0406
11.3661 | -34.5312 | -19.0494
11.3722 | -34.5317 | -19.0581
11.3783 | -34.5321 | -19.0672
11.3844 | -34.5325 | -19.0762
11.3906 | -34.533 | -19.0852
11.3967 | -34.5334 | -19.0939
11.4028 | -34.5338 | -19.1031
11.4089 | -34.5343 | -19.1118
11.415 | -34.5347 | -19.1207
11.4211 | -34.5351 | -19.1298
11.4272 | -34.5355 | -19.1388
11.4333 | -34.5359 | -19.1475
11.4394 | -34.5363 | -19.1565
11.4455 | -34.5367 | -19.1653
11.4516 | -34.5371 | -19.1743
11.4577 | -34.5376 | -19.1833
11.4638 | -34.538 | -19.192
11.4699 | -34.5385 | -19.2011
11.476 | -34.5389 | -19.2104
11.4821 | -34.5393 | -19.2193
11.4882 | -34.5398 | -19.2286
11.4943 | -34.5402 | -19.2377
11.5004 | -34.5406 | -19.2463
11.5065 | -34.5411 | -19.2551
11.5126 | -34.5415 | -19.2639
11.5187 | -34.5419 | -19.273
11.5248 | -34.5423 | -19.2822
11.5309 | -34.5428 | -19.2911
11.5371 | -34.5432 | -19.2999
11.5432 | -34.5437 | -19.3089
11.5493 | -34.5441 | -19.3181
11.5554 | -34.5445 | -19.327
11.5615 | -34.545 | -19.3362
11.5676 | -34.5454 | -19.3452
11.5737 | -34.5458 | -19.3538
11.5798 | -34.5463 | -19.3626
11.5859 | -34.5467 | -19.3715
11.592 | -34.5471 | -19.3803
11.5981 | -34.5476 | -19.3891
11.6042 | -34.548 | -19.3982
11.6103 | -34.5484 | -19.4069
11.6164 | -34.5489 | -19.4157
11.6225 | -34.5493 | -19.4245
11.6286 | -34.5497 | -19.4333
11.6347 | -34.5501 | -19.4423
11.6408 | -34.5506 | -19.4511
11.6469 | -34.551 | -19.46
11.653 | -34.5514 | -19.4687
11.6591 | -34.5519 | -19.4777
11.6652 | -34.5523 | -19.4864
11.6713 | -34.5528 | -19.4952
11.6775 | -34.5532 | -19.504
11.6836 | -34.5536 | -19.5129
11.6897 | -34.5541 | -19.5219
11.6958 | -34.5545 | -19.5309
11.7019 | -34.555 | -19.5399
11.708 | -34.5554 | -19.5488
11.7141 | -34.5559 | -19.5575
11.7202 | -34.5563 | -19.5664
11.7263 | -34.5567 | -19.5754
11.7324 | -34.5571 | -19.5844
11.7385 | -34.5576 | -19.5934
11.7446 | -34.558 | -19.6022
11.7507 | -34.5585 | -19.6108
11.7568 | -34.5589 | -19.6198
11.7629 | -34.5593 | -19.6288
11.769 | -34.5597 | -19.6375
11.7751 | -34.5602 | -19.6463
11.7812 | -34.5607 | -19.6549
11.7873 | -34.5611 | -19.6638
11.7934 | -34.5615 | -19.6728
11.7995 | -34.562 | -19.6815
11.8056 | -34.5624 | -19.6902
11.8117 | -34.5629 | -19.6991
11.8178 | -34.5633 | -19.7078
11.824 | -34.5638 | -19.7165
11.8301 | -34.5642 | -19.7253
11.8362 | -34.5647 | -19.7338
11.8423 | -34.5651 | -19.7426
11.8484 | -34.5655 | -19.7515
11.8545 | -34.566 | -19.7602
11.8606 | -34.5664 | -19.7688
11.8667 | -34.5669 | -19.7778
11.8728 | -34.5673 | -19.7864
11.8789 | -34.5678 | -19.7952
11.885 | -34.5682 | -19.8041
11.8911 | -34.5687 | -19.8128
11.8972 | -34.5691 | -19.8215
11.9033 | -34.5695 | -19.8302
11.9094 | -34.57 | -19.8389
11.9155 | -34.5704 | -19.8476
11.9216 | -34.5709 | -19.8564
11.9277 | -34.5713 | -19.8651
11.9338 | -34.5718 | -19.8738
11.9399 | -34.5722 | -19.8828
11.946 | -34.5727 | -19.8915
11.9521 | -34.5731 | -19.9001
11.9582 | -34.5736 | -19.9089
11.9644 | -34.574 | -19.9174
11.9705 | -34.5745 | -19.9262
11.9766 | -34.575 | -19.935
11.9827 | -34.5754 | -19.9437
11.9888 | -34.5759 | -19.9524
11.9949 | -34.5764 | -19.9608
12.001 | -34.5769 | -19.9697
12.0071 | -34.5773 | -19.9782
12.0132 | -34.5777 | -19.9868
12.0193 | -34.5782 | -19.9956
12.0254 | -34.5786 | -20.0043
12.0315 | -34.5791 | -20.0128
12.0376 | -34.5795 | -20.0213
12.0437 | -34.5799 | -20.03
12.0498 | -34.5804 | -20.0388
12.0559 | -34.5808 | -20.0474
12.062 | -34.5813 | -20.0558
12.0681 | -34.5818 | -20.0646
12.0742 | -34.5822 | -20.0732
12.0803 | -34.5827 | -20.0821
12.0864 | -34.5831 | -20.0909
12.0925 | -34.5836 | -20.0996
12.0986 | -34.5841 | -20.1083
12.1047 | -34.5845 | -20.117
12.1109 | -34.585 | -20.1255
12.117 | -34.5855 | -20.1341
12.1231 | -34.5859 | -20.1428
12.1292 | -34.5864 | -20.1514
12.1353 | -34.5868 | -20.1601
12.1414 | -34.5873 | -20.1689
12.1475 | -34.5877 | -20.1778
12.1536 | -34.5882 | -20.1865
12.1597 | -34.5886 | -20.1952
12.1658 | -34.5891 | -20.2038
12.1719 | -34.5896 | -20.2122
12.178 | -34.59 | -20.2209
12.1841 | -34.5905 | -20.2298
12.1902 | -34.591 | -20.2385
12.1963 | -34.5914 | -20.2473
12.2024 | -34.5919 | -20.2558
12.2085 | -34.5924 | -20.2644
12.2146 | -34.5929 | -20.2732
12.2207 | -34.5933 | -20.2818
12.2268 | -34.5938 | -20.2904
12.2329 | -34.5942 | -20.299
12.239 | -34.5948 | -20.3073
12.2451 | -34.5952 | -20.3159
12.2513 | -34.5957 | -20.3247
12.2574 | -34.5961 | -20.3333
12.2635 | -34.5966 | -20.3419
12.2696 | -34.5971 | -20.3504
12.2757 | -34.5975 | -20.3587
12.2818 | -34.598 | -20.3673
12.2879 | -34.5985 | -20.376
12.294 | -34.599 | -20.3843
12.3001 | -34.5994 | -20.3929
12.3062 | -34.5999 | -20.4016
12.3123 | -34.6004 | -20.4099
12.3184 | -34.6009 | -20.4183
12.3245 | -34.6013 | -20.4267
12.3306 | -34.6018 | -20.4353
12.3367 | -34.6023 | -20.4439
12.3428 | -34.6027 | -20.4525
12.3489 | -34.6032 | -20.4607
12.355 | -34.6037 | -20.4695
12.3611 | -34.6041 | -20.4783
12.3672 | -34.6046 | -20.4867
12.3733 | -34.6051 | -20.4954
12.3794 | -34.6055 | -20.5036
12.3855 | -34.606 | -20.5121
12.3916 | -34.6065 | -20.5206
12.3978 | -34.607 | -20.5294
12.4039 | -34.6074 | -20.5381
12.41 | -34.6079 | -20.5466
12.4161 | -34.6084 | -20.5548
12.4222 | -34.6089 | -20.5634
12.4283 | -34.6093 | -20.5718
12.4344 | -34.6098 | -20.5805
12.4405 | -34.6103 | -20.5892
12.4466 | -34.6107 | -20.5976
12.4527 | -34.6112 | -20.606
12.4588 | -34.6117 | -20.6145
12.4649 | -34.6121 | -20.6231
12.471 | -34.6126 | -20.6318
12.4771 | -34.6131 | -20.6402
12.4832 | -34.6135 | -20.6486
12.4893 | -34.614 | -20.657
12.4954 | -34.6145 | -20.6653
12.5015 | -34.615 | -20.6737
12.5076 | -34.6154 | -20.6823
12.5137 | -34.6159 | -20.6909
12.5198 | -34.6164 | -20.6994
12.5259 | -34.6169 | -20.708
12.532 | -34.6173 | -20.7165
12.5382 | -34.6178 | -20.7251
12.5443 | -34.6182 | -20.7336
12.5504 | -34.6187 | -20.7422
12.5565 | -34.6192 | -20.7506
12.5626 | -34.6197 | -20.7593
12.5687 | -34.6202 | -20.7679
12.5748 | -34.6207 | -20.7764
12.5809 | -34.6211 | -20.7849
12.587 | -34.6216 | -20.7936
12.5931 | -34.6221 | -20.8019
12.5992 | -34.6226 | -20.8104
12.6053 | -34.6231 | -20.8188
12.6114 | -34.6235 | -20.8272
12.6175 | -34.624 | -20.8359
12.6236 | -34.6245 | -20.8444
12.6297 | -34.625 | -20.8528
12.6358 | -34.6254 | -20.861
12.6419 | -34.6259 | -20.8695
12.648 | -34.6264 | -20.8782
12.6541 | -34.6268 | -20.8866
12.6602 | -34.6273 | -20.8946
12.6663 | -34.6278 | -20.9031
12.6724 | -34.6283 | -20.9114
12.6785 | -34.6288 | -20.9197
12.6847 | -34.6293 | -20.9281
12.6908 | -34.6298 | -20.9366
12.6969 | -34.6303 | -20.9451
12.703 | -34.6308 | -20.9536
12.7091 | -34.6312 | -20.962
12.7152 | -34.6317 | -20.9702
12.7213 | -34.6321 | -20.9788
12.7274 | -34.6326 | -20.9873
12.7335 | -34.6331 | -20.9955
12.7396 | -34.6336 | -21.0037
12.7457 | -34.6341 | -21.0123
12.7518 | -34.6345 | -21.0206
12.7579 | -34.635 | -21.0293
12.764 | -34.6355 | -21.0379
12.7701 | -34.6359 | -21.0464
12.7762 | -34.6364 | -21.0548
12.7823 | -34.637 | -21.0633
12.7884 | -34.6375 | -21.0719
12.7945 | -34.6379 | -21.0806
12.8006 | -34.6384 | -21.0889
12.8067 | -34.6389 | -21.0972
12.8128 | -34.6394 | -21.1054
12.8189 | -34.6398 | -21.1139
12.8251 | -34.6403 | -21.1223
12.8312 | -34.6408 | -21.1308
12.8373 | -34.6412 | -21.1392
12.8434 | -34.6417 | -21.1478
12.8495 | -34.6422 | -21.156
12.8556 | -34.6426 | -21.1644
12.8617 | -34.6431 | -21.1728
12.8678 | -34.6436 | -21.1815
12.8739 | -34.6441 | -21.1898
12.88 | -34.6445 | -21.1981
12.8861 | -34.645 | -21.2064
12.8922 | -34.6455 | -21.215
12.8983 | -34.646 | -21.2233
12.9044 | -34.6464 | -21.2314
12.9105 | -34.6469 | -21.2397
12.9166 | -34.6474 | -21.2481
12.9227 | -34.6479 | -21.2565
12.9288 | -34.6483 | -21.265
12.9349 | -34.6488 | -21.2735
12.941 | -34.6493 | -21.282
12.9471 | -34.6498 | -21.2905
12.9532 | -34.6503 | -21.2989
12.9593 | -34.6508 | -21.3075
12.9654 | -34.6512 | -21.316
12.9716 | -34.6517 | -21.3246
12.9777 | -34.6522 | -21.3331
12.9838 | -34.6527 | -21.3413
12.9899 | -34.6532 | -21.3496
12.996 | -34.6536 | -21.358
13.0021 | -34.6541 | -21.3664
13.0082 | -34.6546 | -21.3747
13.0143 | -34.6551 | -21.3828
13.0204 | -34.6556 | -21.391
13.0265 | -34.6561 | -21.3993
13.0326 | -34.6565 | -21.4078
13.0387 | -34.657 | -21.4163
13.0448 | -34.6575 | -21.4247
13.0509 | -34.658 | -21.4331
13.057 | -34.6585 | -21.4414
13.0631 | -34.659 | -21.4497
13.0692 | -34.6595 | -21.4579
13.0753 | -34.66 | -21.4666
13.0814 | -34.6604 | -21.4751
13.0875 | -34.6609 | -21.4834
13.0936 | -34.6614 | -21.4919
13.0997 | -34.6619 | -21.5003
13.1058 | -34.6624 | -21.5086
13.112 | -34.6629 | -21.517
13.1181 | -34.6633 | -21.5252
13.1242 | -34.6638 | -21.5336
13.1303 | -34.6643 | -21.542
13.1364 | -34.6647 | -21.5504
13.1425 | -34.6652 | -21.5585
13.1486 | -34.6656 | -21.5669
13.1547 | -34.6661 | -21.5755
13.1608 | -34.6666 | -21.5838
13.1669 | -34.6671 | -21.592
13.173 | -34.6676 | -21.6004
13.1791 | -34.6681 | -21.6087
13.1852 | -34.6685 | -21.6171
13.1913 | -34.669 | -21.6252
13.1974 | -34.6695 | -21.6337
13.2035 | -34.6699 | -21.642
13.2096 | -34.6704 | -21.6502
13.2157 | -34.6709 | -21.6584
13.2218 | -34.6714 | -21.6666
13.2279 | -34.6719 | -21.6748
13.234 | -34.6724 | -21.6831
13.2401 | -34.6729 | -21.6914
13.2462 | -34.6733 | -21.6997
13.2524 | -34.6738 | -21.7081
13.2585 | -34.6743 | -21.7164
13.2646 | -34.6748 | -21.7247
13.2707 | -34.6753 | -21.7331
13.2768 | -34.6758 | -21.7417
13.2829 | -34.6762 | -21.7502
13.289 | -34.6767 | -21.7583
13.2951 | -34.6772 | -21.7665
13.3012 | -34.6777 | -21.7747
13.3073 | -34.6782 | -21.7833
13.3134 | -34.6786 | -21.7917
13.3195 | -34.6791 | -21.7999
13.3256 | -34.6796 | -21.808
13.3317 | -34.68 | -21.8163
13.3378 | -34.6805 | -21.8246
13.3439 | -34.681 | -21.8332
13.35 | -34.6815 | -21.8417
13.3561 | -34.6819 | -21.8499
13.3622 | -34.6824 | -21.8581
13.3683 | -34.6829 | -21.8663
13.3744 | -34.6833 | -21.8744
13.3805 | -34.6838 | -21.8828
13.3866 | -34.6843 | -21.891
13.3927 | -34.6847 | -21.8994
13.3989 | -34.6852 | -21.9077
13.405 | -34.6856 | -21.916
13.4111 | -34.6861 | -21.924
13.4172 | -34.6866 | -21.9323
13.4233 | -34.687 | -21.9408
13.4294 | -34.6875 | -21.9493
13.4355 | -34.6879 | -21.9574
13.4416 | -34.6884 | -21.9655
13.4477 | -34.6889 | -21.9739
13.4538 | -34.6894 | -21.9822
13.4599 | -34.6898 | -21.9906
13.466 | -34.6903 | -21.999
13.4721 | -34.6907 | -22.0071
13.4782 | -34.6912 | -22.0153
13.4843 | -34.6917 | -22.0238
13.4904 | -34.6922 | -22.0319
13.4965 | -34.6926 | -22.0403
13.5026 | -34.6931 | -22.0487
13.5087 | -34.6936 | -22.057
13.5148 | -34.694 | -22.0656
13.5209 | -34.6945 | -22.074
13.527 | -34.6949 | -22.0822
13.5331 | -34.6954 | -22.0905
13.5393 | -34.6959 | -22.0988
13.5454 | -34.6964 | -22.107
13.5515 | -34.6969 | -22.1155
13.5576 | -34.6973 | -22.1239
13.5637 | -34.6978 | -22.1322
13.5698 | -34.6983 | -22.1406
13.5759 | -34.6988 | -22.1484
13.582 | -34.6993 | -22.1566
13.5881 | -34.6997 | -22.1649
13.5942 | -34.7002 | -22.1735
13.6003 | -34.7007 | -22.1818
13.6064 | -34.7011 | -22.1902
13.6125 | -34.7016 | -22.1988
13.6186 | -34.7021 | -22.2069
13.6247 | -34.7026 | -22.2153
13.6308 | -34.7031 | -22.2237
13.6369 | -34.7035 | -22.232
13.643 | -34.704 | -22.2404
13.6491 | -34.7044 | -22.2487
13.6552 | -34.7049 | -22.257
13.6613 | -34.7054 | -22.2652
13.6674 | -34.7058 | -22.2735
13.6735 | -34.7063 | -22.2818
13.6796 | -34.7068 | -22.2901
13.6858 | -34.7072 | -22.2982
13.6919 | -34.7077 | -22.3064
13.698 | -34.7082 | -22.3147
13.7041 | -34.7086 | -22.3232
13.7102 | -34.7091 | -22.3316
13.7163 | -34.7096 | -22.3398
13.7224 | -34.7101 | -22.3482
13.7285 | -34.7105 | -22.3565
13.7346 | -34.711 | -22.3647
13.7407 | -34.7115 | -22.373
13.7468 | -34.7119 | -22.3814
13.7529 | -34.7124 | -22.3895
13.759 | -34.7128 | -22.3979
13.7651 | -34.7133 | -22.4063
13.7712 | -34.7138 | -22.4146
13.7773 | -34.7142 | -22.423
13.7834 | -34.7147 | -22.4313
13.7895 | -34.7152 | -22.4396
13.7956 | -34.7156 | -22.4481
13.8017 | -34.7161 | -22.4562
13.8078 | -34.7165 | -22.4645
13.8139 | -34.717 | -22.4726
13.82 | -34.7174 | -22.4807
13.8262 | -34.7179 | -22.489
13.8323 | -34.7184 | -22.4975
13.8384 | -34.7188 | -22.5058
13.8445 | -34.7193 | -22.5142
13.8506 | -34.7198 | -22.5225
13.8567 | -34.7202 | -22.5309
13.8628 | -34.7206 | -22.5393
13.8689 | -34.7211 | -22.5477
13.875 | -34.7216 | -22.5559
13.8811 | -34.722 | -22.5641
13.8872 | -34.7224 | -22.5723
13.8933 | -34.7229 | -22.5804
13.8994 | -34.7233 | -22.5887
13.9055 | -34.7238 | -22.5973
13.9116 | -34.7242 | -22.6054
13.9177 | -34.7247 | -22.6138
13.9238 | -34.7251 | -22.6221
13.9299 | -34.7256 | -22.6303
13.936 | -34.726 | -22.6385
13.9421 | -34.7265 | -22.6466
13.9482 | -34.7269 | -22.6548
13.9543 | -34.7274 | -22.6633
13.9604 | -34.7278 | -22.6714
13.9665 | -34.7283 | -22.6796
13.9727 | -34.7287 | -22.6884
13.9788 | -34.7292 | -22.6966
13.9849 | -34.7297 | -22.705
13.991 | -34.7301 | -22.7133
13.9971 | -34.7306 | -22.7214
14.0032 | -34.731 | -22.7297
14.0093 | -34.7315 | -22.738
14.0154 | -34.732 | -22.7461
14.0215 | -34.7324 | -22.7543
14.0276 | -34.7329 | -22.7626
14.0337 | -34.7334 | -22.7709
14.0398 | -34.7338 | -22.779
14.0459 | -34.7343 | -22.7873
14.052 | -34.7348 | -22.7957
14.0581 | -34.7352 | -22.804
14.0642 | -34.7356 | -22.8122
14.0703 | -34.7361 | -22.8205
14.0764 | -34.7366 | -22.8285
14.0825 | -34.737 | -22.8369
14.0886 | -34.7375 | -22.8452
14.0947 | -34.738 | -22.8535
14.1008 | -34.7384 | -22.8617
14.1069 | -34.7389 | -22.8698
14.1131 | -34.7393 | -22.878
14.1192 | -34.7398 | -22.8862
14.1253 | -34.7403 | -22.8942
14.1314 | -34.7407 | -22.9025
14.1375 | -34.7412 | -22.9108
14.1436 | -34.7417 | -22.9191
14.1497 | -34.7422 | -22.9274
14.1558 | -34.7426 | -22.9356
14.1619 | -34.7431 | -22.9438
14.168 | -34.7435 | -22.9521
14.1741 | -34.744 | -22.9603
14.1802 | -34.7445 | -22.9684
14.1863 | -34.7449 | -22.9768
14.1924 | -34.7454 | -22.9851
14.1985 | -34.7459 | -22.9933
14.2046 | -34.7464 | -23.0017
14.2107 | -34.7468 | -23.0097
14.2168 | -34.7473 | -23.018
14.2229 | -34.7477 | -23.0262
14.229 | -34.7482 | -23.0347
14.2351 | -34.7486 | -23.0428
14.2412 | -34.749 | -23.051
14.2473 | -34.7495 | -23.0592
14.2534 | -34.75 | -23.0673
14.2596 | -34.7504 | -23.0753
14.2657 | -34.7508 | -23.0835
14.2718 | -34.7513 | -23.0918
14.2779 | -34.7518 | -23.1
14.284 | -34.7523 | -23.1083
14.2901 | -34.7527 | -23.1165
14.2962 | -34.7532 | -23.1249
14.3023 | -34.7536 | -23.1334
14.3084 | -34.7541 | -23.1415
14.3145 | -34.7545 | -23.1499
14.3206 | -34.755 | -23.1585
14.3267 | -34.7554 | -23.1667
14.3328 | -34.7559 | -23.175
14.3389 | -34.7563 | -23.1832
14.345 | -34.7568 | -23.1914
14.3511 | -34.7572 | -23.1998
14.3572 | -34.7577 | -23.2081
14.3633 | -34.7581 | -23.2167
14.3694 | -34.7586 | -23.2247
14.3755 | -34.7591 | -23.233
14.3816 | -34.7596 | -23.2415
14.3877 | -34.7601 | -23.2496
14.3938 | -34.7605 | -23.2582
14.4 | -34.761 | -23.2665
14.4061 | -34.7615 | -23.2747
14.4122 | -34.7619 | -23.2831
14.4183 | -34.7624 | -23.2911
14.4244 | -34.7629 | -23.2993
14.4305 | -34.7634 | -23.3075
14.4366 | -34.7639 | -23.3156
14.4427 | -34.7643 | -23.324
14.4488 | -34.7648 | -23.3324
14.4549 | -34.7652 | -23.3405
14.461 | -34.7657 | -23.3486
14.4671 | -34.7661 | -23.357
14.4732 | -34.7666 | -23.3653
14.4793 | -34.7671 | -23.3736
14.4854 | -34.7676 | -23.3818
14.4915 | -34.7681 | -23.3901
14.4976 | -34.7685 | -23.398
14.5037 | -34.769 | -23.4061
14.5098 | -34.7695 | -23.4145
14.5159 | -34.7699 | -23.4226
14.522 | -34.7703 | -23.4307
14.5281 | -34.7708 | -23.4387
14.5342 | -34.7713 | -23.4471
14.5403 | -34.7718 | -23.4552
14.5465 | -34.7722 | -23.4635
14.5526 | -34.7726 | -23.4719
14.5587 | -34.7731 | -23.4802
14.5648 | -34.7736 | -23.4884
14.5709 | -34.7741 | -23.4965
14.577 | -34.7745 | -23.5046
14.5831 | -34.7749 | -23.5129
14.5892 | -34.7754 | -23.5211
14.5953 | -34.7759 | -23.5293
14.6014 | -34.7763 | -23.5375
14.6075 | -34.7768 | -23.5456
14.6136 | -34.7772 | -23.5537
14.6197 | -34.7777 | -23.5621
14.6258 | -34.7781 | -23.5702
14.6319 | -34.7786 | -23.5782
14.638 | -34.779 | -23.5865
14.6441 | -34.7794 | -23.5948
14.6502 | -34.7799 | -23.6033
14.6563 | -34.7804 | -23.6115
14.6624 | -34.7808 | -23.6197
14.6685 | -34.7813 | -23.6276
14.6746 | -34.7818 | -23.6358
14.6807 | -34.7822 | -23.6439
14.6869 | -34.7827 | -23.6521
14.693 | -34.7832 | -23.6602
14.6991 | -34.7836 | -23.6684
14.7052 | -34.7841 | -23.6764
14.7113 | -34.7846 | -23.6847
14.7174 | -34.7851 | -23.693
14.7235 | -34.7855 | -23.701
14.7296 | -34.786 | -23.7092
14.7357 | -34.7865 | -23.7173
14.7418 | -34.787 | -23.7255
14.7479 | -34.7874 | -23.7335
14.754 | -34.7879 | -23.7418
14.7601 | -34.7884 | -23.7498
14.7662 | -34.7888 | -23.7579
14.7723 | -34.7893 | -23.7658
14.7784 | -34.7897 | -23.774
14.7845 | -34.7902 | -23.7821
14.7906 | -34.7907 | -23.7905
14.7967 | -34.7911 | -23.7985
14.8028 | -34.7916 | -23.8065
14.8089 | -34.7921 | -23.8147
14.815 | -34.7925 | -23.8229
14.8211 | -34.793 | -23.8311
14.8272 | -34.7934 | -23.8394
14.8334 | -34.7939 | -23.8476
14.8395 | -34.7944 | -23.8557
14.8456 | -34.7949 | -23.8638
14.8517 | -34.7954 | -23.872
14.8578 | -34.7958 | -23.8802
14.8639 | -34.7963 | -23.8885
14.87 | -34.7968 | -23.8968
14.8761 | -34.7973 | -23.9049
14.8822 | -34.7977 | -23.9131
14.8883 | -34.7982 | -23.9213
14.8944 | -34.7986 | -23.9294
14.9005 | -34.7991 | -23.9376
14.9066 | -34.7996 | -23.9457
14.9127 | -34.8001 | -23.9539
14.9188 | -34.8005 | -23.9619
14.9249 | -34.801 | -23.97
14.931 | -34.8015 | -23.9782
14.9371 | -34.8019 | -23.9863
14.9432 | -34.8024 | -23.9944
14.9493 | -34.8029 | -24.0025
14.9554 | -34.8034 | -24.0107
14.9615 | -34.8038 | -24.0188
14.9676 | -34.8043 | -24.0271
14.9738 | -34.8048 | -24.035
14.9799 | -34.8052 | -24.0432
14.986 | -34.8057 | -24.0512
14.9921 | -34.8062 | -24.0593
14.9982 | -34.8066 | -24.0674
15.0043 | -34.8071 | -24.0753
15.0104 | -34.8075 | -24.0836
15.0165 | -34.808 | -24.092
15.0226 | -34.8085 | -24.1
15.0287 | -34.809 | -24.1082
15.0348 | -34.8095 | -24.1163
15.0409 | -34.8099 | -24.1244
15.047 | -34.8104 | -24.1326
15.0531 | -34.8109 | -24.1406
15.0592 | -34.8114 | -24.1489
15.0653 | -34.8119 | -24.157
15.0714 | -34.8123 | -24.1649
15.0775 | -34.8128 | -24.1729
15.0836 | -34.8133 | -24.181
15.0897 | -34.8138 | -24.189
15.0958 | -34.8143 | -24.1971
15.1019 | -34.8148 | -24.2053
15.108 | -34.8153 | -24.2133
15.1141 | -34.8157 | -24.2214
15.1203 | -34.8162 | -24.2296
15.1264 | -34.8167 | -24.2377
15.1325 | -34.8172 | -24.2458
15.1386 | -34.8177 | -24.2539
15.1447 | -34.8181 | -24.262
15.1508 | -34.8187 | -24.2701
15.1569 | -34.8191 | -24.2781
15.163 | -34.8197 | -24.2859
15.1691 | -34.8201 | -24.294
15.1752 | -34.8206 | -24.3019
15.1813 | -34.8211 | -24.31
15.1874 | -34.8216 | -24.3181
15.1935 | -34.8221 | -24.326
15.1996 | -34.8226 | -24.3339
15.2057 | -34.823 | -24.3419
15.2118 | -34.8235 | -24.3499
15.2179 | -34.824 | -24.3582
15.224 | -34.8245 | -24.3662
15.2301 | -34.825 | -24.3741
15.2362 | -34.8255 | -24.3822
15.2423 | -34.8259 | -24.3905
15.2484 | -34.8264 | -24.3986
15.2545 | -34.8269 | -24.4068
15.2607 | -34.8274 | -24.4149
15.2668 | -34.828 | -24.4226
15.2729 | -34.8285 | -24.4306
15.279 | -34.8289 | -24.4389
15.2851 | -34.8294 | -24.4469
15.2912 | -34.8299 | -24.455
15.2973 | -34.8304 | -24.4632
15.3034 | -34.8309 | -24.4711
15.3095 | -34.8314 | -24.4793
15.3156 | -34.8319 | -24.4872
15.3217 | -34.8324 | -24.4952
15.3278 | -34.8329 | -24.5035
15.3339 | -34.8334 | -24.5113
15.34 | -34.8339 | -24.5192
15.3461 | -34.8344 | -24.5273
15.3522 | -34.8349 | -24.5352
15.3583 | -34.8354 | -24.5435
15.3644 | -34.8359 | -24.5513
15.3705 | -34.8364 | -24.559
15.3766 | -34.8369 | -24.5671
15.3827 | -34.8374 | -24.5749
15.3888 | -34.8379 | -24.583
15.3949 | -34.8384 | -24.5911
15.401 | -34.8389 | -24.5992
15.4072 | -34.8394 | -24.6072
15.4133 | -34.8399 | -24.6151
15.4194 | -34.8404 | -24.6233
15.4255 | -34.8408 | -24.6313
15.4316 | -34.8413 | -24.6393
15.4377 | -34.8418 | -24.6475
15.4438 | -34.8423 | -24.6554
15.4499 | -34.8428 | -24.6634
15.456 | -34.8433 | -24.6713
15.4621 | -34.8438 | -24.6792
15.4682 | -34.8443 | -24.6874
15.4743 | -34.8449 | -24.6953
15.4804 | -34.8453 | -24.7033
15.4865 | -34.8458 | -24.7113
15.4926 | -34.8463 | -24.7194
15.4987 | -34.8468 | -24.7275
15.5048 | -34.8473 | -24.7356
15.5109 | -34.8478 | -24.7435
15.517 | -34.8483 | -24.7516
15.5231 | -34.8488 | -24.7596
15.5292 | -34.8493 | -24.7676
15.5353 | -34.8498 | -24.7757
15.5414 | -34.8503 | -24.7836
15.5476 | -34.8508 | -24.7912
15.5537 | -34.8513 | -24.7993
15.5598 | -34.8518 | -24.8075
15.5659 | -34.8523 | -24.8155
15.572 | -34.8529 | -24.8235
15.5781 | -34.8533 | -24.8314
15.5842 | -34.8539 | -24.8391
15.5903 | -34.8544 | -24.8468
15.5964 | -34.8549 | -24.8547
15.6025 | -34.8554 | -24.8625
15.6086 | -34.8559 | -24.8706
15.6147 | -34.8564 | -24.8786
15.6208 | -34.857 | -24.8864
15.6269 | -34.8575 | -24.8942
15.633 | -34.858 | -24.9022
15.6391 | -34.8585 | -24.9103
15.6452 | -34.859 | -24.9184
15.6513 | -34.8596 | -24.9262
15.6574 | -34.8601 | -24.9344
15.6635 | -34.8606 | -24.9422
15.6696 | -34.8611 | -24.9503
15.6757 | -34.8616 | -24.9583
15.6818 | -34.8621 | -24.9663
15.688 | -34.8626 | -24.9741
15.6941 | -34.8631 | -24.9819
15.7002 | -34.8636 | -24.9898
15.7063 | -34.8641 | -24.9975
15.7124 | -34.8646 | -25.0056
15.7185 | -34.8651 | -25.0135
15.7246 | -34.8656 | -25.0214
15.7307 | -34.8662 | -25.0293
15.7368 | -34.8667 | -25.0373
15.7429 | -34.8672 | -25.0451
15.749 | -34.8677 | -25.0531
15.7551 | -34.8682 | -25.0612
15.7612 | -34.8688 | -25.0691
15.7673 | -34.8693 | -25.077
15.7734 | -34.8698 | -25.0848
15.7795 | -34.8703 | -25.0927
15.7856 | -34.8709 | -25.1004
15.7917 | -34.8714 | -25.1082
15.7978 | -34.8719 | -25.1163
15.8039 | -34.8724 | -25.1243
15.81 | -34.8729 | -25.1322
15.8161 | -34.8734 | -25.1401
15.8222 | -34.874 | -25.1479
15.8283 | -34.8745 | -25.1556
15.8345 | -34.8751 | -25.1636
15.8406 | -34.8756 | -25.1714
15.8467 | -34.8761 | -25.1793
15.8528 | -34.8766 | -25.1873
15.8589 | -34.8771 | -25.1951
15.865 | -34.8777 | -25.2028
15.8711 | -34.8782 | -25.2106
15.8772 | -34.8788 | -25.2183
15.8833 | -34.8793 | -25.2263
15.8894 | -34.8798 | -25.2345
15.8955 | -34.8803 | -25.2422
15.9016 | -34.8808 | -25.2497
15.9077 | -34.8813 | -25.2578
15.9138 | -34.8818 | -25.2656
15.9199 | -34.8824 | -25.2733
15.926 | -34.8829 | -25.281
15.9321 | -34.8834 | -25.2888
15.9382 | -34.8839 | -25.2966
15.9443 | -34.8845 | -25.3046
15.9504 | -34.885 | -25.3126
15.9565 | -34.8855 | -25.3203
15.9626 | -34.886 | -25.3284
15.9687 | -34.8865 | -25.336
15.9749 | -34.887 | -25.3437
15.981 | -34.8875 | -25.3514
15.9871 | -34.888 | -25.3594
15.9932 | -34.8885 | -25.3671
15.9993 | -34.8891 | -25.375
16.0054 | -34.8896 | -25.3824
16.0115 | -34.8902 | -25.3903
16.0176 | -34.8907 | -25.3982
16.0237 | -34.8912 | -25.4062
16.0298 | -34.8917 | -25.4139
16.0359 | -34.8922 | -25.4216
16.042 | -34.8928 | -25.4293
16.0481 | -34.8933 | -25.4372
16.0542 | -34.8939 | -25.4452
16.0603 | -34.8944 | -25.4529
16.0664 | -34.8949 | -25.4606
16.0725 | -34.8954 | -25.4679
16.0786 | -34.8959 | -25.4756
16.0847 | -34.8964 | -25.4836
16.0908 | -34.8969 | -25.4915
16.0969 | -34.8974 | -25.4994
16.103 | -34.8979 | -25.5072
16.1091 | -34.8985 | -25.5148
16.1152 | -34.899 | -25.5225
16.1214 | -34.8995 | -25.5303
16.1275 | -34.9001 | -25.538
16.1336 | -34.9006 | -25.5458
16.1397 | -34.9012 | -25.5537
16.1458 | -34.9017 | -25.5612
16.1519 | -34.9023 | -25.5692
16.158 | -34.9028 | -25.5771
16.1641 | -34.9032 | -25.5849
16.1702 | -34.9038 | -25.5927
16.1763 | -34.9043 | -25.6005
16.1824 | -34.9048 | -25.608
16.1885 | -34.9054 | -25.6156
16.1946 | -34.9059 | -25.6232
16.2007 | -34.9064 | -25.631
16.2068 | -34.907 | -25.6388
16.2129 | -34.9075 | -25.6467
16.219 | -34.908 | -25.6541
16.2251 | -34.9086 | -25.6617
16.2312 | -34.9091 | -25.6694
16.2373 | -34.9097 | -25.677
16.2434 | -34.9102 | -25.6846
16.2495 | -34.9108 | -25.6922
16.2556 | -34.9113 | -25.7001
16.2618 | -34.9118 | -25.7078
16.2679 | -34.9124 | -25.7156
16.274 | -34.9129 | -25.7234
16.2801 | -34.9134 | -25.731
16.2862 | -34.9139 | -25.7387
16.2923 | -34.9145 | -25.7466
16.2984 | -34.915 | -25.7543
16.3045 | -34.9155 | -25.762
16.3106 | -34.916 | -25.7697
16.3167 | -34.9165 | -25.7774
16.3228 | -34.9171 | -25.785
16.3289 | -34.9175 | -25.7927
16.335 | -34.9181 | -25.8003
16.3411 | -34.9186 | -25.8079
16.3472 | -34.9192 | -25.8159
16.3533 | -34.9197 | -25.8236
16.3594 | -34.9203 | -25.8309
16.3655 | -34.9208 | -25.8387
16.3716 | -34.9213 | -25.8461
16.3777 | -34.9218 | -25.8538
16.3838 | -34.9223 | -25.8615
16.3899 | -34.9229 | -25.8691
16.396 | -34.9234 | -25.8769
16.4021 | -34.9239 | -25.8845
16.4083 | -34.9245 | -25.892
16.4144 | -34.925 | -25.8996
16.4205 | -34.9255 | -25.9073
16.4266 | -34.9261 | -25.9147
16.4327 | -34.9266 | -25.9223
16.4388 | -34.9271 | -25.9299
16.4449 | -34.9276 | -25.9376
16.451 | -34.9281 | -25.9454
16.4571 | -34.9287 | -25.9533
16.4632 | -34.9292 | -25.9605
16.4693 | -34.9297 | -25.9682
16.4754 | -34.9302 | -25.976
16.4815 | -34.9307 | -25.9837
16.4876 | -34.9312 | -25.9911
16.4937 | -34.9318 | -25.9986
16.4998 | -34.9323 | -26.0059
16.5059 | -34.9328 | -26.0134
16.512 | -34.9333 | -26.0212
16.5181 | -34.9339 | -26.0292
16.5242 | -34.9344 | -26.037
16.5303 | -34.9349 | -26.0445
16.5364 | -34.9354 | -26.052
16.5425 | -34.936 | -26.0596
16.5487 | -34.9365 | -26.067
16.5548 | -34.937 | -26.0748
16.5609 | -34.9375 | -26.0825
16.567 | -34.9381 | -26.09
16.5731 | -34.9386 | -26.0978
16.5792 | -34.9391 | -26.1054
16.5853 | -34.9397 | -26.113
16.5914 | -34.9402 | -26.1206
16.5975 | -34.9407 | -26.1283
16.6036 | -34.9413 | -26.1356
16.6097 | -34.9418 | -26.1432
16.6158 | -34.9424 | -26.1506
16.6219 | -34.9429 | -26.1584
16.628 | -34.9435 | -26.166
16.6341 | -34.944 | -26.1735
16.6402 | -34.9445 | -26.1807
16.6463 | -34.945 | -26.1883
16.6524 | -34.9455 | -26.1957
16.6585 | -34.946 | -26.2034
16.6646 | -34.9466 | -26.2109
16.6707 | -34.9471 | -26.2182
16.6768 | -34.9477 | -26.2259
16.6829 | -34.9482 | -26.2334
16.689 | -34.9487 | -26.2411
16.6952 | -34.9492 | -26.2485
16.7013 | -34.9498 | -26.2562
16.7074 | -34.9503 | -26.2639
16.7135 | -34.9508 | -26.2716
16.7196 | -34.9514 | -26.2791
16.7257 | -34.952 | -26.2869
16.7318 | -34.9525 | -26.2943
16.7379 | -34.9531 | -26.3018
16.744 | -34.9536 | -26.3092
16.7501 | -34.9541 | -26.317
16.7562 | -34.9546 | -26.3246
16.7623 | -34.9551 | -26.3321
16.7684 | -34.9556 | -26.3397
16.7745 | -34.9561 | -26.3475
16.7806 | -34.9567 | -26.3545
16.7867 | -34.9572 | -26.3621
16.7928 | -34.9578 | -26.3698
16.7989 | -34.9583 | -26.3772
16.805 | -34.9588 | -26.3847
16.8111 | -34.9593 | -26.3921
16.8172 | -34.9598 | -26.3996
16.8233 | -34.9604 | -26.4073
16.8294 | -34.9609 | -26.415
16.8356 | -34.9614 | -26.4227
16.8417 | -34.962 | -26.4302
16.8478 | -34.9625 | -26.4377
16.8539 | -34.963 | -26.4451
16.86 | -34.9635 | -26.4529
16.8661 | -34.9641 | -26.4604
16.8722 | -34.9646 | -26.4681
16.8783 | -34.9651 | -26.4758
16.8844 | -34.9657 | -26.4831
16.8905 | -34.9662 | -26.4906
16.8966 | -34.9667 | -26.4979
16.9027 | -34.9673 | -26.5056
16.9088 | -34.9678 | -26.5134
16.9149 | -34.9683 | -26.5209
16.921 | -34.9689 | -26.5282
16.9271 | -34.9694 | -26.5358
16.9332 | -34.97 | -26.5433
16.9393 | -34.9705 | -26.5508
16.9454 | -34.971 | -26.5581
16.9515 | -34.9715 | -26.5657
16.9576 | -34.972 | -26.573
16.9637 | -34.9725 | -26.5805
16.9698 | -34.9731 | -26.588
16.9759 | -34.9736 | -26.5953
16.9821 | -34.9742 | -26.603
16.9882 | -34.9747 | -26.6102
16.9943 | -34.9752 | -26.6175
17.0004 | -34.9758 | -26.6249
17.0065 | -34.9763 | -26.6325
17.0126 | -34.9768 | -26.64
17.0187 | -34.9773 | -26.6477
17.0248 | -34.9779 | -26.6547
17.0309 | -34.9784 | -26.6621
17.037 | -34.9789 | -26.6696
17.0431 | -34.9794 | -26.6771
17.0492 | -34.9799 | -26.6847
17.0553 | -34.9804 | -26.692
17.0614 | -34.981 | -26.6996
17.0675 | -34.9815 | -26.7072
17.0736 | -34.982 | -26.7147
17.0797 | -34.9825 | -26.7219
17.0858 | -34.983 | -26.7297
17.0919 | -34.9835 | -26.7373
17.098 | -34.984 | -26.7445
17.1041 | -34.9846 | -26.7519
17.1102 | -34.9851 | -26.7594
17.1163 | -34.9856 | -26.7668
17.1225 | -34.9862 | -26.7741
17.1286 | -34.9867 | -26.7813
17.1347 | -34.9873 | -26.7889
17.1408 | -34.9878 | -26.7966
17.1469 | -34.9883 | -26.8038
17.153 | -34.9889 | -26.8113
17.1591 | -34.9894 | -26.8188
17.1652 | -34.9899 | -26.8261
17.1713 | -34.9905 | -26.8334
17.1774 | -34.991 | -26.8408
17.1835 | -34.9915 | -26.848
17.1896 | -34.992 | -26.8553
17.1957 | -34.9926 | -26.8628
17.2018 | -34.9931 | -26.8697
17.2079 | -34.9936 | -26.8773
17.214 | -34.9942 | -26.8848
17.2201 | -34.9947 | -26.8922
17.2262 | -34.9952 | -26.8998
17.2323 | -34.9957 | -26.9073
17.2384 | -34.9963 | -26.9148
17.2445 | -34.9968 | -26.9221
17.2506 | -34.9973 | -26.9294
17.2567 | -34.9978 | -26.9369
17.2628 | -34.9983 | -26.9441
17.269 | -34.9988 | -26.9514
17.2751 | -34.9994 | -26.9588
17.2812 | -34.9998 | -26.9662
17.2873 | -35.0004 | -26.9735
17.2934 | -35.0009 | -26.981
17.2995 | -35.0014 | -26.9887
17.3056 | -35.0019 | -26.9959
17.3117 | -35.0025 | -27.0034
17.3178 | -35.003 | -27.0107
17.3239 | -35.0035 | -27.0182
17.33 | -35.0039 | -27.0256
17.3361 | -35.0044 | -27.033
17.3422 | -35.0049 | -27.0403
17.3483 | -35.0055 | -27.0477
17.3544 | -35.006 | -27.0552
17.3605 | -35.0065 | -27.0625
17.3666 | -35.007 | -27.0699
17.3727 | -35.0075 | -27.0774
17.3788 | -35.0081 | -27.0846
17.3849 | -35.0086 | -27.0919
17.391 | -35.0091 | -27.0992
17.3971 | -35.0096 | -27.1064
17.4032 | -35.0101 | -27.1139
17.4094 | -35.0106 | -27.1213
17.4155 | -35.0112 | -27.1286
17.4216 | -35.0117 | -27.136
17.4277 | -35.0122 | -27.1436
17.4338 | -35.0127 | -27.151
17.4399 | -35.0133 | -27.1584
17.446 | -35.0138 | -27.1656
17.4521 | -35.0143 | -27.1731
17.4582 | -35.0148 | -27.1803
17.4643 | -35.0153 | -27.1877
17.4704 | -35.0158 | -27.1952
17.4765 | -35.0163 | -27.2028
17.4826 | -35.0168 | -27.21
17.4887 | -35.0173 | -27.2175
17.4948 | -35.0178 | -27.2248
17.5009 | -35.0184 | -27.2321
17.507 | -35.0189 | -27.2397
17.5131 | -35.0194 | -27.2471
17.5192 | -35.0199 | -27.2541
17.5253 | -35.0204 | -27.2614
17.5314 | -35.0209 | -27.2687
17.5375 | -35.0214 | -27.2761
17.5436 | -35.0219 | -27.2837
17.5497 | -35.0224 | -27.291
17.5559 | -35.0229 | -27.298
17.562 | -35.0234 | -27.3052
17.5681 | -35.0239 | -27.3126
17.5742 | -35.0244 | -27.32
17.5803 | -35.025 | -27.3273
17.5864 | -35.0255 | -27.3345
17.5925 | -35.026 | -27.342
17.5986 | -35.0265 | -27.3491
17.6047 | -35.027 | -27.3565
17.6108 | -35.0275 | -27.3638
17.6169 | -35.028 | -27.371
17.623 | -35.0285 | -27.3781
17.6291 | -35.029 | -27.3853
17.6352 | -35.0295 | -27.3927
17.6413 | -35.03 | -27.4
17.6474 | -35.0305 | -27.4075
17.6535 | -35.031 | -27.4147
17.6596 | -35.0315 | -27.4219
17.6657 | -35.0321 | -27.4294
17.6718 | -35.0326 | -27.4367
17.6779 | -35.0331 | -27.4438
17.684 | -35.0336 | -27.4511
17.6901 | -35.0341 | -27.4584
17.6963 | -35.0346 | -27.4652
17.7024 | -35.0351 | -27.4725
17.7085 | -35.0356 | -27.4798
17.7146 | -35.0361 | -27.487
17.7207 | -35.0367 | -27.4941
17.7268 | -35.0372 | -27.5015
17.7329 | -35.0377 | -27.5089
17.739 | -35.0382 | -27.5161
17.7451 | -35.0387 | -27.5232
17.7512 | -35.0392 | -27.5305
17.7573 | -35.0397 | -27.538
17.7634 | -35.0402 | -27.545
17.7695 | -35.0407 | -27.5522
17.7756 | -35.0412 | -27.5595
17.7817 | -35.0417 | -27.5668
17.7878 | -35.0422 | -27.5742
17.7939 | -35.0427 | -27.5816
17.8 | -35.0432 | -27.5888
17.8061 | -35.0437 | -27.596
17.8122 | -35.0443 | -27.6033
17.8183 | -35.0448 | -27.6107
17.8244 | -35.0453 | -27.6181
17.8305 | -35.0457 | -27.6253
17.8366 | -35.0463 | -27.6322
17.8428 | -35.0468 | -27.6397
17.8489 | -35.0473 | -27.6469
17.855 | -35.0478 | -27.654
17.8611 | -35.0482 | -27.661
17.8672 | -35.0487 | -27.6681
17.8733 | -35.0492 | -27.6755
17.8794 | -35.0498 | -27.6826
17.8855 | -35.0503 | -27.69
17.8916 | -35.0508 | -27.6973
17.8977 | -35.0513 | -27.7046
17.9038 | -35.0518 | -27.7116
17.9099 | -35.0523 | -27.7187
17.916 | -35.0528 | -27.7259
17.9221 | -35.0533 | -27.7331
17.9282 | -35.0538 | -27.7404
17.9343 | -35.0544 | -27.7476
17.9404 | -35.0548 | -27.7546
17.9465 | -35.0553 | -27.762
17.9526 | -35.0558 | -27.7694
17.9587 | -35.0563 | -27.7767
17.9648 | -35.0568 | -27.7839
17.9709 | -35.0573 | -27.7912
17.977 | -35.0578 | -27.7985
17.9832 | -35.0583 | -27.8056
17.9893 | -35.0588 | -27.8131
17.9954 | -35.0593 | -27.8204
18.0015 | -35.0597 | -27.8277
18.0076 | -35.0602 | -27.835
18.0137 | -35.0607 | -27.8421
18.0198 | -35.0611 | -27.849
18.0259 | -35.0616 | -27.8563
18.032 | -35.0621 | -27.8635
18.0381 | -35.0626 | -27.8708
18.0442 | -35.0631 | -27.8781
18.0503 | -35.0636 | -27.8853
18.0564 | -35.064 | -27.8925
18.0625 | -35.0646 | -27.8999
18.0686 | -35.065 | -27.9073
18.0747 | -35.0655 | -27.9146
18.0808 | -35.066 | -27.9216
18.0869 | -35.0665 | -27.9288
18.093 | -35.067 | -27.936
18.0991 | -35.0675 | -27.943
18.1052 | -35.068 | -27.9504
18.1113 | -35.0684 | -27.9576
18.1174 | -35.0689 | -27.9647
18.1236 | -35.0694 | -27.972
18.1297 | -35.0699 | -27.9792
18.1358 | -35.0704 | -27.9864
18.1419 | -35.071 | -27.9935
18.148 | -35.0715 | -28.0007
18.1541 | -35.072 | -28.0079
18.1602 | -35.0725 | -28.0152
18.1663 | -35.0729 | -28.0222
18.1724 | -35.0734 | -28.0294
18.1785 | -35.0739 | -28.0365
18.1846 | -35.0744 | -28.0434
18.1907 | -35.0748 | -28.0506
18.1968 | -35.0753 | -28.0577
18.2029 | -35.0758 | -28.0649
18.209 | -35.0763 | -28.0722
18.2151 | -35.0768 | -28.0795
18.2212 | -35.0773 | -28.0869
18.2273 | -35.0778 | -28.0941
18.2334 | -35.0783 | -28.1013
18.2395 | -35.0788 | -28.1087
18.2456 | -35.0792 | -28.1158
18.2517 | -35.0797 | -28.123
18.2578 | -35.0802 | -28.1299
18.2639 | -35.0807 | -28.1373
18.2701 | -35.0812 | -28.1444
18.2762 | -35.0817 | -28.1517
18.2823 | -35.0822 | -28.1586
18.2884 | -35.0827 | -28.1657
18.2945 | -35.0832 | -28.1726
18.3006 | -35.0837 | -28.1797
18.3067 | -35.0842 | -28.1869
18.3128 | -35.0846 | -28.1941
18.3189 | -35.0851 | -28.2014
18.325 | -35.0856 | -28.2082
18.3311 | -35.0861 | -28.2153
18.3372 | -35.0866 | -28.2225
18.3433 | -35.0871 | -28.2297
18.3494 | -35.0875 | -28.237
18.3555 | -35.088 | -28.2442
18.3616 | -35.0885 | -28.2513
18.3677 | -35.089 | -28.2586
18.3738 | -35.0895 | -28.2657
18.3799 | -35.0899 | -28.2728
18.386 | -35.0904 | -28.2799
18.3921 | -35.0909 | -28.2869
18.3982 | -35.0913 | -28.2937
18.4043 | -35.0918 | -28.3007
18.4105 | -35.0922 | -28.308
18.4166 | -35.0927 | -28.3152
18.4227 | -35.0932 | -28.3224
18.4288 | -35.0937 | -28.3293
18.4349 | -35.0942 | -28.3364
18.441 | -35.0946 | -28.3433
18.4471 | -35.0951 | -28.3504
18.4532 | -35.0956 | -28.3577
18.4593 | -35.0961 | -28.3648
18.4654 | -35.0966 | -28.3718
18.4715 | -35.097 | -28.3791
18.4776 | -35.0975 | -28.3864
18.4837 | -35.098 | -28.3937
18.4898 | -35.0984 | -28.4007
18.4959 | -35.0989 | -28.4079
18.502 | -35.0994 | -28.415
18.5081 | -35.0998 | -28.4223
18.5142 | -35.1003 | -28.4295
18.5203 | -35.1008 | -28.4366
18.5264 | -35.1013 | -28.4437
18.5325 | -35.1018 | -28.4511
18.5386 | -35.1022 | -28.4587
18.5447 | -35.1027 | -28.4658
18.5508 | -35.1033 | -28.4727
18.557 | -35.1037 | -28.4803
18.5631 | -35.1042 | -28.4874
18.5692 | -35.1047 | -28.4944
18.5753 | -35.1052 | -28.5008
18.5814 | -35.1057 | -28.5077
18.5875 | -35.1061 | -28.515
18.5936 | -35.1066 | -28.5218
18.5997 | -35.107 | -28.5288
18.6058 | -35.1075 | -28.5359
18.6119 | -35.108 | -28.5431
18.618 | -35.1085 | -28.5502
18.6241 | -35.1089 | -28.5573
18.6302 | -35.1093 | -28.5643
18.6363 | -35.1098 | -28.5715
18.6424 | -35.1103 | -28.5787
18.6485 | -35.1107 | -28.5855
18.6546 | -35.1112 | -28.5925
18.6607 | -35.1116 | -28.5995
18.6668 | -35.1122 | -28.6064
18.6729 | -35.1126 | -28.6139
18.679 | -35.1131 | -28.6209
18.6851 | -35.1136 | -28.628
18.6912 | -35.114 | -28.635
18.6974 | -35.1145 | -28.6417
18.7035 | -35.1149 | -28.6489
18.7096 | -35.1154 | -28.6565
18.7157 | -35.1158 | -28.6634
18.7218 | -35.1163 | -28.6701
18.7279 | -35.1168 | -28.6771
18.734 | -35.1173 | -28.6842
18.7401 | -35.1177 | -28.6914
18.7462 | -35.1182 | -28.6985
18.7523 | -35.1187 | -28.7058
18.7584 | -35.1192 | -28.7129
18.7645 | -35.1197 | -28.7198
18.7706 | -35.1202 | -28.727
18.7767 | -35.1206 | -28.7339
18.7828 | -35.1211 | -28.741
18.7889 | -35.1216 | -28.7478
18.795 | -35.122 | -28.755
18.8011 | -35.1225 | -28.7618
18.8072 | -35.1229 | -28.7691
18.8133 | -35.1234 | -28.7761
18.8194 | -35.1239 | -28.783
18.8255 | -35.1243 | -28.7898
18.8316 | -35.1248 | -28.7969
18.8377 | -35.1252 | -28.8038
18.8439 | -35.1257 | -28.8109
18.85 | -35.1262 | -28.8182
18.8561 | -35.1267 | -28.8253
18.8622 | -35.1271 | -28.8322
18.8683 | -35.1276 | -28.8393
18.8744 | -35.1281 | -28.8464
18.8805 | -35.1285 | -28.8533
18.8866 | -35.1289 | -28.8605
18.8927 | -35.1294 | -28.8673
18.8988 | -35.1299 | -28.8743
18.9049 | -35.1303 | -28.8815
18.911 | -35.1308 | -28.8885
18.9171 | -35.1313 | -28.8956
18.9232 | -35.1317 | -28.9025
18.9293 | -35.1322 | -28.9097
18.9354 | -35.1327 | -28.9167
18.9415 | -35.1332 | -28.9235
18.9476 | -35.1337 | -28.9302
18.9537 | -35.1341 | -28.9374
18.9598 | -35.1346 | -28.9441
18.9659 | -35.1351 | -28.9512
18.972 | -35.1356 | -28.9582
18.9781 | -35.136 | -28.9651
18.9843 | -35.1364 | -28.9723
18.9904 | -35.1369 | -28.9797
18.9965 | -35.1374 | -28.9864
19.0026 | -35.1379 | -28.9938
19.0087 | -35.1383 | -29.0007
19.0148 | -35.1388 | -29.0075
19.0209 | -35.1392 | -29.0146
19.027 | -35.1398 | -29.0214
19.0331 | -35.1402 | -29.0287
19.0392 | -35.1407 | -29.0358
19.0453 | -35.1412 | -29.0428
19.0514 | -35.1416 | -29.0498
19.0575 | -35.142 | -29.0571
19.0636 | -35.1426 | -29.0641
19.0697 | -35.143 | -29.0712
19.0758 | -35.1436 | -29.078
19.0819 | -35.1441 | -29.0851
19.088 | -35.1445 | -29.0923
19.0941 | -35.145 | -29.0996
19.1002 | -35.1454 | -29.1068
19.1063 | -35.1459 | -29.1137
19.1124 | -35.1463 | -29.1203
19.1185 | -35.1468 | -29.1271
19.1246 | -35.1473 | -29.1343
19.1308 | -35.1478 | -29.1415
19.1369 | -35.1483 | -29.1488
19.143 | -35.1487 | -29.1554
19.1491 | -35.1492 | -29.1622
19.1552 | -35.1496 | -29.1692
19.1613 | -35.1501 | -29.176
19.1674 | -35.1505 | -29.1829
19.1735 | -35.151 | -29.1898
19.1796 | -35.1515 | -29.1965
19.1857 | -35.152 | -29.2034
19.1918 | -35.1525 | -29.211
19.1979 | -35.1529 | -29.2182
19.204 | -35.1534 | -29.225
19.2101 | -35.1539 | -29.2318
19.2162 | -35.1543 | -29.2389
19.2223 | -35.1548 | -29.2457
19.2284 | -35.1551 | -29.2525
19.2345 | -35.1556 | -29.2595
19.2406 | -35.1561 | -29.2664
19.2467 | -35.1566 | -29.2733
19.2528 | -35.1571 | -29.2807
19.2589 | -35.1576 | -29.2872
19.265 | -35.158 | -29.2944
19.2712 | -35.1585 | -29.301
19.2773 | -35.159 | -29.3079
19.2834 | -35.1595 | -29.3147
19.2895 | -35.16 | -29.3217
19.2956 | -35.1604 | -29.3288
19.3017 | -35.1608 | -29.336
19.3078 | -35.1613 | -29.3429
19.3139 | -35.1618 | -29.3494
19.32 | -35.1623 | -29.3562
19.3261 | -35.1628 | -29.3632
19.3322 | -35.1632 | -29.3704
19.3383 | -35.1637 | -29.3766
19.3444 | -35.1641 | -29.3837
19.3505 | -35.1646 | -29.3903
19.3566 | -35.1652 | -29.3972
19.3627 | -35.1656 | -29.4045
19.3688 | -35.1662 | -29.4111
19.3749 | -35.1667 | -29.4183
19.381 | -35.1672 | -29.4251
19.3871 | -35.1676 | -29.4319
19.3932 | -35.1681 | -29.4391
19.3993 | -35.1687 | -29.4463
19.4054 | -35.1691 | -29.4534
19.4115 | -35.1696 | -29.4605
19.4177 | -35.17 | -29.4671
19.4238 | -35.1706 | -29.4739
19.4299 | -35.171 | -29.481
19.436 | -35.1715 | -29.4884
19.4421 | -35.1719 | -29.4953
19.4482 | -35.1724 | -29.5027
19.4543 | -35.1728 | -29.5094
19.4604 | -35.1732 | -29.5163
19.4665 | -35.1736 | -29.5237
19.4726 | -35.174 | -29.5305
19.4787 | -35.1745 | -29.5376
19.4848 | -35.175 | -29.5444
19.4909 | -35.1753 | -29.5515
19.497 | -35.1758 | -29.5586
19.5031 | -35.1762 | -29.5652
19.5092 | -35.1766 | -29.5723
19.5153 | -35.1771 | -29.5793
19.5214 | -35.1776 | -29.5865
19.5275 | -35.178 | -29.5929
19.5336 | -35.1785 | -29.5996
19.5397 | -35.1789 | -29.6066
19.5458 | -35.1794 | -29.6129
19.5519 | -35.1799 | -29.6197
19.5581 | -35.1804 | -29.6263
19.5642 | -35.1808 | -29.6332
19.5703 | -35.1812 | -29.6403
19.5764 | -35.1817 | -29.6476
19.5825 | -35.1822 | -29.6541
19.5886 | -35.1827 | -29.6613
19.5947 | -35.1832 | -29.6679
19.6008 | -35.1836 | -29.6748
19.6069 | -35.1841 | -29.6818
19.613 | -35.1846 | -29.6884
19.6191 | -35.1851 | -29.6952
19.6252 | -35.1855 | -29.7023
19.6313 | -35.186 | -29.7084
19.6374 | -35.1864 | -29.7149
19.6435 | -35.1868 | -29.7216
19.6496 | -35.1873 | -29.7289
19.6557 | -35.1878 | -29.7356
19.6618 | -35.1881 | -29.7416
19.6679 | -35.1886 | -29.7486
19.674 | -35.189 | -29.7554
19.6801 | -35.1894 | -29.7624
19.6862 | -35.19 | -29.7696
19.6923 | -35.1904 | -29.7765
19.6984 | -35.1909 | -29.7833
19.7046 | -35.1914 | -29.7901
19.7107 | -35.1919 | -29.7965
19.7168 | -35.1924 | -29.8035
19.7229 | -35.1928 | -29.8103
19.729 | -35.1933 | -29.8172
19.7351 | -35.1938 | -29.8239
19.7412 | -35.1943 | -29.8307
19.7473 | -35.1947 | -29.838
19.7534 | -35.1952 | -29.8458
19.7595 | -35.1957 | -29.8526
19.7656 | -35.1961 | -29.86
19.7717 | -35.1965 | -29.8656
19.7778 | -35.1969 | -29.8729
19.7839 | -35.1974 | -29.88
19.79 | -35.1978 | -29.8867
19.7961 | -35.1982 | -29.8934
19.8022 | -35.1986 | -29.8998
19.8083 | -35.199 | -29.9067
19.8144 | -35.1996 | -29.9133
19.8205 | -35.2 | -29.9201
19.8266 | -35.2004 | -29.9272
19.8327 | -35.2009 | -29.9337
19.8388 | -35.2013 | -29.9411
19.845 | -35.2018 | -29.948
19.8511 | -35.2022 | -29.9552
19.8572 | -35.2026 | -29.9623
19.8633 | -35.2031 | -29.9693
19.8694 | -35.2034 | -29.976
19.8755 | -35.2038 | -29.9834
19.8816 | -35.2044 | -29.9909
19.8877 | -35.2048 | -29.9978
19.8938 | -35.2053 | -30.0046
19.8999 | -35.2057 | -30.0121
19.906 | -35.2061 | -30.0184
19.9121 | -35.2066 | -30.0252
19.9182 | -35.207 | -30.0316
19.9243 | -35.2075 | -30.0384
19.9304 | -35.2079 | -30.0462
19.9365 | -35.2084 | -30.0528
19.9426 | -35.209 | -30.0598
19.9487 | -35.2095 | -30.0665
19.9548 | -35.21 | -30.0734
19.9609 | -35.2104 | -30.08
19.967 | -35.2108 | -30.0872
19.9731 | -35.2113 | -30.0937
19.9792 | -35.2118 | -30.1002
19.9853 | -35.2122 | -30.1071
19.9915 | -35.2127 | -30.1142
19.9976 | -35.2131 | -30.1208
20.0037 | -35.2135 | -30.1279
20.0098 | -35.2139 | -30.1345
20.0159 | -35.2144 | -30.1416
20.022 | -35.2148 | -30.1489
20.0281 | -35.2153 | -30.1552
20.0342 | -35.2158 | -30.1622
20.0403 | -35.2164 | -30.169
20.0464 | -35.2168 | -30.1763
20.0525 | -35.2174 | -30.1824
20.0586 | -35.2179 | -30.1896
20.0647 | -35.2183 | -30.1963
20.0708 | -35.2188 | -30.2029
20.0769 | -35.2193 | -30.2095
20.083 | -35.2197 | -30.216
20.0891 | -35.22 | -30.223
20.0952 | -35.2204 | -30.2297
20.1013 | -35.2209 | -30.2367
20.1074 | -35.2214 | -30.2435
20.1135 | -35.2218 | -30.2503
20.1196 | -35.2223 | -30.2564
20.1257 | -35.2227 | -30.2631
20.1319 | -35.2232 | -30.2696
20.138 | -35.2236 | -30.2765
20.1441 | -35.2241 | -30.284
20.1502 | -35.2245 | -30.2913
20.1563 | -35.225 | -30.2984
20.1624 | -35.2254 | -30.3049
20.1685 | -35.2258 | -30.3125
20.1746 | -35.2263 | -30.3189
20.1807 | -35.2267 | -30.3255
20.1868 | -35.2271 | -30.3323
20.1929 | -35.2276 | -30.3394
20.199 | -35.2281 | -30.3462
20.2051 | -35.2285 | -30.3527
20.2112 | -35.229 | -30.3596
20.2173 | -35.2294 | -30.3661
20.2234 | -35.2299 | -30.3729
20.2295 | -35.2304 | -30.3798
20.2356 | -35.2308 | -30.3871
20.2417 | -35.2313 | -30.394
20.2478 | -35.2318 | -30.4008
20.2539 | -35.2322 | -30.408
20.26 | -35.2327 | -30.4146
20.2661 | -35.2331 | -30.4214
20.2723 | -35.2336 | -30.4285
20.2784 | -35.2342 | -30.4351
20.2845 | -35.2346 | -30.4421
20.2906 | -35.235 | -30.4489
20.2967 | -35.2354 | -30.4555
20.3028 | -35.2359 | -30.462
20.3089 | -35.2364 | -30.4686
20.315 | -35.2368 | -30.4749
20.3211 | -35.2373 | -30.4819
20.3272 | -35.2377 | -30.4888
20.3333 | -35.2383 | -30.4953
20.3394 | -35.2387 | -30.5022
20.3455 | -35.2392 | -30.5085
20.3516 | -35.2397 | -30.5154
20.3577 | -35.2402 | -30.5222
20.3638 | -35.2407 | -30.5289
20.3699 | -35.2411 | -30.5354
20.376 | -35.2416 | -30.5422
20.3821 | -35.2421 | -30.549
20.3882 | -35.2425 | -30.5558
20.3943 | -35.243 | -30.5626
20.4004 | -35.2435 | -30.5695
20.4065 | -35.244 | -30.5762
20.4126 | -35.2444 | -30.5825
20.4188 | -35.2449 | -30.5893
20.4249 | -35.2454 | -30.596
20.431 | -35.2458 | -30.6028
20.4371 | -35.2462 | -30.6091
20.4432 | -35.2466 | -30.6159
20.4493 | -35.2471 | -30.6225
20.4554 | -35.2475 | -30.6291
20.4615 | -35.248 | -30.6359
20.4676 | -35.2485 | -30.643
20.4737 | -35.2489 | -30.6503
20.4798 | -35.2494 | -30.6569
20.4859 | -35.2498 | -30.6637
20.492 | -35.2503 | -30.6703
20.4981 | -35.2508 | -30.6771
20.5042 | -35.2512 | -30.684
20.5103 | -35.2516 | -30.6907
20.5164 | -35.2521 | -30.6973
20.5225 | -35.2526 | -30.7038
20.5286 | -35.2531 | -30.71
20.5347 | -35.2535 | -30.7164
20.5408 | -35.254 | -30.7228
20.5469 | -35.2545 | -30.7299
20.553 | -35.2549 | -30.7367
20.5592 | -35.2554 | -30.7431
20.5653 | -35.2558 | -30.7494
20.5714 | -35.2564 | -30.7561
20.5775 | -35.2569 | -30.7628
20.5836 | -35.2574 | -30.7697
20.5897 | -35.2579 | -30.7766
20.5958 | -35.2582 | -30.7834
20.6019 | -35.2587 | -30.7904
20.608 | -35.2592 | -30.7972
20.6141 | -35.2597 | -30.8036
20.6202 | -35.2601 | -30.8105
20.6263 | -35.2606 | -30.8173
20.6324 | -35.2611 | -30.8235
20.6385 | -35.2616 | -30.8303
20.6446 | -35.2621 | -30.8372
20.6507 | -35.2625 | -30.844
20.6568 | -35.2629 | -30.8504
20.6629 | -35.2634 | -30.857
20.669 | -35.2638 | -30.864
20.6751 | -35.2642 | -30.8709
20.6812 | -35.2647 | -30.8772
20.6873 | -35.2651 | -30.884
20.6934 | -35.2655 | -30.8906
20.6995 | -35.266 | -30.8969
20.7057 | -35.2665 | -30.9036
20.7118 | -35.267 | -30.9107
20.7179 | -35.2675 | -30.9171
20.724 | -35.2679 | -30.9237
20.7301 | -35.2684 | -30.9304
20.7362 | -35.2688 | -30.9374
20.7423 | -35.2693 | -30.9441
20.7484 | -35.2697 | -30.9509
20.7545 | -35.2703 | -30.9573
20.7606 | -35.2707 | -30.9642
20.7667 | -35.2712 | -30.9706
20.7728 | -35.2717 | -30.9769
20.7789 | -35.2722 | -30.9834
20.785 | -35.2726 | -30.9901
20.7911 | -35.2731 | -30.9964
20.7972 | -35.2735 | -31.0034
20.8033 | -35.2739 | -31.0103
20.8094 | -35.2744 | -31.0169
20.8155 | -35.2748 | -31.0238
20.8216 | -35.2752 | -31.0307
20.8277 | -35.2757 | -31.0369
20.8338 | -35.2762 | -31.0434
20.8399 | -35.2767 | -31.0503
20.8461 | -35.2772 | -31.0573
20.8522 | -35.2777 | -31.0643
20.8583 | -35.2782 | -31.0708
20.8644 | -35.2787 | -31.0777
20.8705 | -35.2792 | -31.0848
20.8766 | -35.2796 | -31.0915
20.8827 | -35.2801 | -31.0985
20.8888 | -35.2805 | -31.1054
20.8949 | -35.281 | -31.1116
20.901 | -35.2815 | -31.1184
20.9071 | -35.2819 | -31.1249
20.9132 | -35.2824 | -31.1315
20.9193 | -35.2828 | -31.1384
20.9254 | -35.2833 | -31.1454
20.9315 | -35.2838 | -31.1518
20.9376 | -35.2843 | -31.1583
20.9437 | -35.2847 | -31.165
20.9498 | -35.2852 | -31.1716
20.9559 | -35.2857 | -31.1783
20.962 | -35.2862 | -31.1853
20.9681 | -35.2866 | -31.1924
20.9742 | -35.2871 | -31.199
20.9803 | -35.2875 | -31.2056
20.9864 | -35.288 | -31.2123
20.9926 | -35.2885 | -31.2191
20.9987 | -35.289 | -31.2253
21.0048 | -35.2896 | -31.2317
21.0109 | -35.29 | -31.239
21.017 | -35.2905 | -31.2454
21.0231 | -35.2911 | -31.2521
21.0292 | -35.2915 | -31.2589
21.0353 | -35.2919 | -31.2654
21.0414 | -35.2924 | -31.2724
21.0475 | -35.2929 | -31.2791
21.0536 | -35.2933 | -31.2857
21.0597 | -35.2937 | -31.292
21.0658 | -35.2942 | -31.2983
21.0719 | -35.2947 | -31.305
21.078 | -35.2951 | -31.3118
21.0841 | -35.2956 | -31.319
21.0902 | -35.296 | -31.3258
21.0963 | -35.2965 | -31.3325
21.1024 | -35.2969 | -31.3393
21.1085 | -35.2973 | -31.346
21.1146 | -35.2977 | -31.3525
21.1207 | -35.2981 | -31.3588
21.1268 | -35.2985 | -31.3655
21.133 | -35.299 | -31.3722
21.1391 | -35.2996 | -31.3786
21.1452 | -35.2999 | -31.3853
21.1513 | -35.3004 | -31.3922
21.1574 | -35.3008 | -31.3992
21.1635 | -35.3013 | -31.4058
21.1696 | -35.3018 | -31.4121
21.1757 | -35.3022 | -31.4186
21.1818 | -35.3027 | -31.4256
21.1879 | -35.3032 | -31.432
21.194 | -35.3037 | -31.4384
21.2001 | -35.3042 | -31.4452
21.2062 | -35.3047 | -31.4514
21.2123 | -35.3051 | -31.4577
21.2184 | -35.3056 | -31.4645
21.2245 | -35.3062 | -31.4704
21.2306 | -35.3065 | -31.4756
21.2367 | -35.3069 | -31.4825
21.2428 | -35.3074 | -31.4885
21.2489 | -35.3079 | -31.4953
21.255 | -35.3084 | -31.502
21.2611 | -35.3087 | -31.5089
21.2672 | -35.3092 | -31.5159
21.2733 | -35.3097 | -31.5222
21.2795 | -35.31 | -31.5292
21.2856 | -35.3105 | -31.5364
21.2917 | -35.3109 | -31.5432
21.2978 | -35.3115 | -31.5496
21.3039 | -35.312 | -31.5562
21.31 | -35.3125 | -31.5626
21.3161 | -35.3131 | -31.5686
21.3222 | -35.3136 | -31.5763
21.3283 | -35.314 | -31.5824
21.3344 | -35.3144 | -31.5885
21.3405 | -35.3149 | -31.5956
21.3466 | -35.3154 | -31.602
21.3527 | -35.3158 | -31.6079
21.3588 | -35.3163 | -31.6145
21.3649 | -35.3168 | -31.6208
21.371 | -35.3173 | -31.6274
21.3771 | -35.3178 | -31.6338
21.3832 | -35.3182 | -31.6398
21.3893 | -35.3187 | -31.6473
21.3954 | -35.3191 | -31.6539
21.4015 | -35.3196 | -31.6601
21.4076 | -35.3201 | -31.6668
21.4137 | -35.3206 | -31.6739
21.4199 | -35.3211 | -31.6807
21.426 | -35.3216 | -31.6865
21.4321 | -35.3219 | -31.6929
21.4382 | -35.3223 | -31.6994
21.4443 | -35.3227 | -31.7056
21.4504 | -35.3232 | -31.712
21.4565 | -35.3237 | -31.7182
21.4626 | -35.3241 | -31.725
21.4687 | -35.3246 | -31.7317
21.4748 | -35.325 | -31.7382
21.4809 | -35.3255 | -31.7446
21.487 | -35.3259 | -31.7516
21.4931 | -35.3264 | -31.7582
21.4992 | -35.3269 | -31.7652
21.5053 | -35.3273 | -31.7724
21.5114 | -35.3279 | -31.7797
21.5175 | -35.3283 | -31.7861
21.5236 | -35.3287 | -31.7927
21.5297 | -35.3291 | -31.799
21.5358 | -35.3296 | -31.8062
21.5419 | -35.3301 | -31.8126
21.548 | -35.3305 | -31.8195
21.5541 | -35.3309 | -31.8262
21.5602 | -35.3313 | -31.8323
21.5664 | -35.3318 | -31.839
21.5725 | -35.3323 | -31.8459
21.5786 | -35.3327 | -31.8526
21.5847 | -35.3332 | -31.8593
21.5908 | -35.3337 | -31.866
21.5969 | -35.334 | -31.8724
21.603 | -35.3345 | -31.8793
21.6091 | -35.335 | -31.886
21.6152 | -35.3354 | -31.8931
21.6213 | -35.3359 | -31.8995
21.6274 | -35.3363 | -31.9062
21.6335 | -35.3367 | -31.9126
21.6396 | -35.3372 | -31.9189
21.6457 | -35.3377 | -31.9252
21.6518 | -35.3382 | -31.9314
21.6579 | -35.3386 | -31.9379
21.664 | -35.3391 | -31.9446
21.6701 | -35.3396 | -31.9512
21.6762 | -35.34 | -31.9579
21.6823 | -35.3405 | -31.9652
21.6884 | -35.341 | -31.9715
21.6945 | -35.3414 | -31.9776
21.7006 | -35.3419 | -31.9844
21.7068 | -35.3423 | -31.9905
21.7129 | -35.3428 | -31.9966
21.719 | -35.3433 | -32.0037
21.7251 | -35.3438 | -32.0108
21.7312 | -35.3443 | -32.017
21.7373 | -35.3447 | -32.023
21.7434 | -35.3451 | -32.029
21.7495 | -35.3456 | -32.0362
21.7556 | -35.346 | -32.0428
21.7617 | -35.3464 | -32.0494
21.7678 | -35.3469 | -32.0559
21.7739 | -35.3474 | -32.0621
21.78 | -35.3478 | -32.0678
21.7861 | -35.3483 | -32.0746
21.7922 | -35.3488 | -32.0807
21.7983 | -35.3492 | -32.0869
21.8044 | -35.3497 | -32.0936
21.8105 | -35.3501 | -32.1003
21.8166 | -35.3506 | -32.1063
21.8227 | -35.351 | -32.1123
21.8288 | -35.3514 | -32.1186
21.8349 | -35.3518 | -32.1251
21.841 | -35.3523 | -32.1316
21.8471 | -35.3527 | -32.1377
21.8533 | -35.3532 | -32.144
21.8594 | -35.3537 | -32.1502
21.8655 | -35.3542 | -32.157
21.8716 | -35.3545 | -32.1635
21.8777 | -35.355 | -32.1702
21.8838 | -35.3555 | -32.1762
21.8899 | -35.3559 | -32.1827
21.896 | -35.3563 | -32.1897
21.9021 | -35.3569 | -32.1969
21.9082 | -35.3573 | -32.2032
21.9143 | -35.3576 | -32.2092
21.9204 | -35.358 | -32.2159
21.9265 | -35.3585 | -32.2226
21.9326 | -35.359 | -32.2292
21.9387 | -35.3594 | -32.2362
21.9448 | -35.3599 | -32.2431
21.9509 | -35.3604 | -32.25
21.957 | -35.3609 | -32.2568
21.9631 | -35.3613 | -32.2636
21.9692 | -35.3617 | -32.2701
21.9753 | -35.3621 | -32.2759
21.9814 | -35.3626 | -32.2831
21.9875 | -35.3629 | -32.2892
21.9937 | -35.3633 | -32.2953
21.9998 | -35.3639 | -32.3014
22.0059 | -35.3642 | -32.3077
22.012 | -35.3647 | -32.3136
22.0181 | -35.3653 | -32.3203
22.0242 | -35.3656 | -32.3268
22.0303 | -35.366 | -32.3332
22.0364 | -35.3665 | -32.339
22.0425 | -35.3668 | -32.3456
22.0486 | -35.3673 | -32.3514
22.0547 | -35.3678 | -32.3578
22.0608 | -35.3681 | -32.3642
22.0669 | -35.3686 | -32.3707
22.073 | -35.3691 | -32.3766
22.0791 | -35.3695 | -32.3831
22.0852 | -35.37 | -32.389
22.0913 | -35.3704 | -32.3948
22.0974 | -35.3708 | -32.4001
22.1035 | -35.3713 | -32.4071
22.1096 | -35.3717 | -32.4136
22.1157 | -35.3722 | -32.4201
22.1218 | -35.3726 | -32.4266
22.1279 | -35.3732 | -32.4323
22.134 | -35.3736 | -32.4391
22.1402 | -35.3741 | -32.4442
22.1463 | -35.3746 | -32.4513
22.1524 | -35.3752 | -32.4579
22.1585 | -35.3757 | -32.4643
22.1646 | -35.3762 | -32.4712
22.1707 | -35.3766 | -32.4784
22.1768 | -35.3771 | -32.4845
22.1829 | -35.3776 | -32.4913
22.189 | -35.378 | -32.4974
22.1951 | -35.3782 | -32.5032
22.2012 | -35.3787 | -32.5088
22.2073 | -35.3792 | -32.5147
22.2134 | -35.3796 | -32.5212
22.2195 | -35.3802 | -32.5274
22.2256 | -35.3806 | -32.5341
22.2317 | -35.3811 | -32.5406
22.2378 | -35.3816 | -32.5473
22.2439 | -35.382 | -32.5537
22.25 | -35.3824 | -32.5598
22.2561 | -35.3828 | -32.5662
22.2622 | -35.3833 | -32.5723
22.2683 | -35.3837 | -32.5785
22.2744 | -35.3841 | -32.5846
22.2806 | -35.3844 | -32.5913
22.2867 | -35.3848 | -32.5975
22.2928 | -35.3853 | -32.6039
22.2989 | -35.3857 | -32.6101
22.305 | -35.3862 | -32.616
22.3111 | -35.3866 | -32.6225
22.3172 | -35.387 | -32.629
22.3233 | -35.3875 | -32.6354
22.3294 | -35.3881 | -32.6418
22.3355 | -35.3884 | -32.6476
22.3416 | -35.3889 | -32.6537
22.3477 | -35.3893 | -32.66
22.3538 | -35.3896 | -32.6656
22.3599 | -35.39 | -32.6719
22.366 | -35.3904 | -32.6783
22.3721 | -35.3909 | -32.6851
22.3782 | -35.3913 | -32.6909
22.3843 | -35.3917 | -32.697
22.3904 | -35.3922 | -32.703
22.3965 | -35.3927 | -32.7095
22.4026 | -35.3931 | -32.7157
22.4087 | -35.3935 | -32.7218
22.4148 | -35.3939 | -32.728
22.4209 | -35.3944 | -32.734
22.4271 | -35.3948 | -32.74
22.4332 | -35.3952 | -32.746
22.4393 | -35.3955 | -32.7522
22.4454 | -35.3959 | -32.7587
22.4515 | -35.3964 | -32.765
22.4576 | -35.3969 | -32.7714
22.4637 | -35.3972 | -32.778
22.4698 | -35.3977 | -32.7839
22.4759 | -35.3981 | -32.7903
22.482 | -35.3986 | -32.7969
22.4881 | -35.399 | -32.803
22.4942 | -35.3994 | -32.8091
22.5003 | -35.3998 | -32.816
22.5064 | -35.4003 | -32.8225
22.5125 | -35.4007 | -32.8284
22.5186 | -35.4012 | -32.8347
22.5247 | -35.4017 | -32.8411
22.5308 | -35.4021 | -32.8474
22.5369 | -35.4024 | -32.8539
22.543 | -35.4029 | -32.8607
22.5491 | -35.4034 | -32.8669
22.5552 | -35.4038 | -32.8732
22.5613 | -35.4043 | -32.8793
22.5675 | -35.4047 | -32.8852
22.5736 | -35.405 | -32.8915
22.5797 | -35.4055 | -32.8981
22.5858 | -35.406 | -32.9034
22.5919 | -35.4064 | -32.9092
22.598 | -35.4068 | -32.9158
22.6041 | -35.4072 | -32.9219
22.6102 | -35.4076 | -32.9285
22.6163 | -35.4081 | -32.9342
22.6224 | -35.4086 | -32.941
22.6285 | -35.409 | -32.9465
22.6346 | -35.4095 | -32.953
22.6407 | -35.4099 | -32.9593
22.6468 | -35.4103 | -32.9651
22.6529 | -35.4107 | -32.9711
22.659 | -35.4112 | -32.9773
22.6651 | -35.4116 | -32.9838
22.6712 | -35.412 | -32.9903
22.6773 | -35.4124 | -32.9962
22.6834 | -35.4129 | -33.0027
22.6895 | -35.4132 | -33.0075
22.6956 | -35.4135 | -33.0138
22.7017 | -35.4139 | -33.0199
22.7079 | -35.4144 | -33.0263
22.714 | -35.4149 | -33.0323
22.7201 | -35.4153 | -33.0387
22.7262 | -35.4158 | -33.0447
22.7323 | -35.4162 | -33.0507
22.7384 | -35.4166 | -33.0569
22.7445 | -35.417 | -33.0632
22.7506 | -35.4174 | -33.0693
22.7567 | -35.4177 | -33.0759
22.7628 | -35.4181 | -33.0825
22.7689 | -35.4185 | -33.0887
22.775 | -35.419 | -33.0953
22.7811 | -35.4195 | -33.102
22.7872 | -35.42 | -33.1086
22.7933 | -35.4204 | -33.1147
22.7994 | -35.4209 | -33.1209
22.8055 | -35.4213 | -33.1268
22.8116 | -35.4217 | -33.1332
22.8177 | -35.4221 | -33.1394
22.8238 | -35.4226 | -33.1455
22.8299 | -35.423 | -33.1517
22.836 | -35.4234 | -33.1578
22.8421 | -35.4238 | -33.1642
22.8482 | -35.4243 | -33.1705
22.8544 | -35.4248 | -33.1768
22.8605 | -35.4252 | -33.1828
22.8666 | -35.4256 | -33.1884
22.8727 | -35.426 | -33.1945
22.8788 | -35.4264 | -33.2004
22.8849 | -35.4268 | -33.2065
22.891 | -35.4273 | -33.2129
22.8971 | -35.4277 | -33.2192
22.9032 | -35.4281 | -33.2249
22.9093 | -35.4285 | -33.2313
22.9154 | -35.429 | -33.2376
22.9215 | -35.4295 | -33.244
22.9276 | -35.4298 | -33.2501
22.9337 | -35.4302 | -33.2556
22.9398 | -35.4307 | -33.2619
22.9459 | -35.4311 | -33.2679
22.952 | -35.4315 | -33.2738
22.9581 | -35.432 | -33.2798
22.9642 | -35.4324 | -33.2858
22.9703 | -35.4328 | -33.2917
22.9764 | -35.4331 | -33.2976
22.9825 | -35.4336 | -33.3034
22.9886 | -35.4339 | -33.3094
22.9948 | -35.4344 | -33.3153
23.0009 | -35.4347 | -33.3213
23.007 | -35.4352 | -33.3264
23.0131 | -35.4357 | -33.3328
23.0192 | -35.4361 | -33.3382
23.0253 | -35.4366 | -33.344
23.0314 | -35.437 | -33.3499
23.0375 | -35.4375 | -33.3564
23.0436 | -35.438 | -33.3617
23.0497 | -35.4384 | -33.3679
23.0558 | -35.4388 | -33.3735
23.0619 | -35.4392 | -33.3793
23.068 | -35.4396 | -33.3853
23.0741 | -35.44 | -33.3907
23.0802 | -35.4404 | -33.397
23.0863 | -35.4408 | -33.4033
23.0924 | -35.4412 | -33.4093
23.0985 | -35.4416 | -33.4156
23.1046 | -35.442 | -33.4217
23.1107 | -35.4423 | -33.4275
23.1168 | -35.4428 | -33.4337
23.1229 | -35.4432 | -33.4399
23.129 | -35.4437 | -33.4456
23.1351 | -35.4441 | -33.4513
23.1413 | -35.4446 | -33.4579
23.1474 | -35.445 | -33.4629
23.1535 | -35.4454 | -33.4684
23.1596 | -35.4458 | -33.4749
23.1657 | -35.4463 | -33.4814
23.1718 | -35.4467 | -33.4871
23.1779 | -35.4472 | -33.4933
23.184 | -35.4476 | -33.4985
23.1901 | -35.4481 | -33.5042
23.1962 | -35.4484 | -33.5101
23.2023 | -35.4488 | -33.5159
23.2084 | -35.4491 | -33.522
23.2145 | -35.4496 | -33.5279
23.2206 | -35.45 | -33.5334
23.2267 | -35.4504 | -33.5397
23.2328 | -35.4508 | -33.5456
23.2389 | -35.4513 | -33.5513
23.245 | -35.4517 | -33.5569
23.2511 | -35.4521 | -33.563
23.2572 | -35.4525 | -33.569
23.2633 | -35.4529 | -33.5752
23.2694 | -35.4533 | -33.5813
23.2755 | -35.4537 | -33.5873
23.2817 | -35.4541 | -33.5926
23.2878 | -35.4546 | -33.5987
23.2939 | -35.455 | -33.6049
23.3 | -35.4555 | -33.6107
23.3061 | -35.4559 | -33.6171
23.3122 | -35.4563 | -33.6226
23.3183 | -35.4567 | -33.6291
23.3244 | -35.4571 | -33.635
23.3305 | -35.4576 | -33.6405
23.3366 | -35.4579 | -33.6469
23.3427 | -35.4582 | -33.6529
23.3488 | -35.4586 | -33.6589
23.3549 | -35.4591 | -33.6648
23.361 | -35.4594 | -33.6704
23.3671 | -35.4598 | -33.6768
23.3732 | -35.4603 | -33.6823
23.3793 | -35.4608 | -33.6884
23.3854 | -35.4611 | -33.6948
23.3915 | -35.4615 | -33.7006
23.3976 | -35.4619 | -33.7068
23.4037 | -35.4623 | -33.7128
23.4098 | -35.4627 | -33.719
23.4159 | -35.4631 | -33.7249
23.422 | -35.4636 | -33.7314
23.4282 | -35.464 | -33.7362
23.4343 | -35.4645 | -33.7423
23.4404 | -35.4649 | -33.7487
23.4465 | -35.4653 | -33.7549
23.4526 | -35.4657 | -33.7612
23.4587 | -35.4662 | -33.7675
23.4648 | -35.4666 | -33.773
23.4709 | -35.467 | -33.7787
23.477 | -35.4674 | -33.7846
23.4831 | -35.4677 | -33.7906
23.4892 | -35.4682 | -33.7965
23.4953 | -35.4685 | -33.8026
23.5014 | -35.469 | -33.8081
23.5075 | -35.4694 | -33.8144
23.5136 | -35.4698 | -33.8208
23.5197 | -35.4703 | -33.8268
23.5258 | -35.4707 | -33.8329
23.5319 | -35.4711 | -33.8378
23.538 | -35.4715 | -33.8438
23.5441 | -35.4718 | -33.8497
23.5502 | -35.4723 | -33.855
23.5563 | -35.4728 | -33.8606
23.5624 | -35.4731 | -33.8664
23.5686 | -35.4735 | -33.8729
23.5747 | -35.474 | -33.8793
23.5808 | -35.4744 | -33.8854
23.5869 | -35.4748 | -33.891
23.593 | -35.4751 | -33.8967
23.5991 | -35.4756 | -33.9025
23.6052 | -35.476 | -33.9084
23.6113 | -35.4764 | -33.9149
23.6174 | -35.4768 | -33.92
23.6235 | -35.4771 | -33.9257
23.6296 | -35.4775 | -33.9318
23.6357 | -35.4779 | -33.9379
23.6418 | -35.4782 | -33.9438
23.6479 | -35.4786 | -33.95
23.654 | -35.4791 | -33.9557
23.6601 | -35.4795 | -33.9618
23.6662 | -35.4799 | -33.9671
23.6723 | -35.4802 | -33.9726
23.6784 | -35.4807 | -33.9786
23.6845 | -35.4811 | -33.9845
23.6906 | -35.4816 | -33.9904
23.6967 | -35.482 | -33.9958
23.7028 | -35.4824 | -34.0016
23.7089 | -35.4828 | -34.0067
23.7151 | -35.4832 | -34.0126
23.7212 | -35.4836 | -34.0186
23.7273 | -35.4839 | -34.0244
23.7334 | -35.4844 | -34.0302
23.7395 | -35.4848 | -34.036
23.7456 | -35.4852 | -34.0411
23.7517 | -35.4856 | -34.0463
23.7578 | -35.486 | -34.052
23.7639 | -35.4864 | -34.058
23.77 | -35.4868 | -34.0638
23.7761 | -35.4871 | -34.0697
23.7822 | -35.4876 | -34.0754
23.7883 | -35.488 | -34.0811
23.7944 | -35.4885 | -34.0867
23.8005 | -35.4889 | -34.0923
23.8066 | -35.4893 | -34.0986
23.8127 | -35.4897 | -34.1045
23.8188 | -35.49 | -34.1103
23.8249 | -35.4905 | -34.1163
23.831 | -35.4908 | -34.1217
23.8371 | -35.4913 | -34.1276
23.8432 | -35.4917 | -34.1336
23.8493 | -35.4921 | -34.1389
23.8555 | -35.4924 | -34.1448
23.8616 | -35.4928 | -34.1501
23.8677 | -35.4932 | -34.1559
23.8738 | -35.4936 | -34.1616
23.8799 | -35.494 | -34.1679
23.886 | -35.4943 | -34.173
23.8921 | -35.4946 | -34.1779
23.8982 | -35.495 | -34.1842
23.9043 | -35.4954 | -34.1899
23.9104 | -35.4958 | -34.1949
23.9165 | -35.4962 | -34.2006
23.9226 | -35.4966 | -34.2059
23.9287 | -35.497 | -34.2119
23.9348 | -35.4974 | -34.2181
23.9409 | -35.4979 | -34.2242
23.947 | -35.4983 | -34.2302
23.9531 | -35.4988 | -34.2349
23.9592 | -35.4991 | -34.2406
23.9653 | -35.4995 | -34.2467
23.9714 | -35.4998 | -34.2524
23.9775 | -35.5002 | -34.258
23.9836 | -35.5006 | -34.2635
23.9897 | -35.501 | -34.2687
23.9958 | -35.5014 | -34.2743
24.002 | -35.5018 | -34.2802
24.0081 | -35.5021 | -34.2857
24.0142 | -35.5025 | -34.2909
24.0203 | -35.5029 | -34.2964
24.0264 | -35.5032 | -34.3026
24.0325 | -35.5037 | -34.3086
24.0386 | -35.504 | -34.3146
24.0447 | -35.5044 | -34.32
24.0508 | -35.5048 | -34.3253
24.0569 | -35.5051 | -34.3309
24.063 | -35.5056 | -34.336
24.0691 | -35.506 | -34.3416
24.0752 | -35.5064 | -34.3477
24.0813 | -35.5068 | -34.3531
24.0874 | -35.5071 | -34.3589
24.0935 | -35.5075 | -34.3645
24.0996 | -35.5078 | -34.3696
24.1057 | -35.5082 | -34.3754
24.1118 | -35.5086 | -34.3811
24.1179 | -35.509 | -34.3874
24.124 | -35.5095 | -34.3935
24.1301 | -35.5098 | -34.3986
24.1362 | -35.5101 | -34.4043
24.1424 | -35.5105 | -34.41
24.1485 | -35.5108 | -34.416
24.1546 | -35.5112 | -34.4215
24.1607 | -35.5116 | -34.4273
24.1668 | -35.512 | -34.4322
24.1729 | -35.5124 | -34.4379
24.179 | -35.5128 | -34.4433
24.1851 | -35.5132 | -34.4491
24.1912 | -35.5136 | -34.4547
24.1973 | -35.514 | -34.4603
24.2034 | -35.5144 | -34.4663
24.2095 | -35.5148 | -34.4721
24.2156 | -35.5152 | -34.4779
24.2217 | -35.5156 | -34.4837
24.2278 | -35.516 | -34.4893
24.2339 | -35.5164 | -34.4948
24.24 | -35.5168 | -34.5002
24.2461 | -35.5172 | -34.5056
24.2522 | -35.5176 | -34.5114
24.2583 | -35.518 | -34.5171
24.2644 | -35.5184 | -34.523
24.2705 | -35.5188 | -34.5284
24.2766 | -35.5192 | -34.5337
24.2827 | -35.5196 | -34.5395
24.2889 | -35.5199 | -34.5451
24.295 | -35.5203 | -34.5508
24.3011 | -35.5206 | -34.5561
24.3072 | -35.5209 | -34.5608
24.3133 | -35.5213 | -34.5666
24.3194 | -35.5217 | -34.5726
24.3255 | -35.522 | -34.5782
24.3316 | -35.5223 | -34.5839
24.3377 | -35.5227 | -34.5894
24.3438 | -35.5231 | -34.5948
24.3499 | -35.5235 | -34.6004
24.356 | -35.5238 | -34.6056
24.3621 | -35.5241 | -34.6111
24.3682 | -35.5245 | -34.6167
24.3743 | -35.5248 | -34.6226
24.3804 | -35.5252 | -34.6281
24.3865 | -35.5255 | -34.6337
24.3926 | -35.5259 | -34.6393
24.3987 | -35.5263 | -34.6451
24.4048 | -35.5267 | -34.6506
24.4109 | -35.527 | -34.6559
24.417 | -35.5274 | -34.6619
24.4231 | -35.5277 | -34.6672
24.4293 | -35.5281 | -34.6725
24.4354 | -35.5285 | -34.6781
24.4415 | -35.5289 | -34.6833
24.4476 | -35.5293 | -34.6883
24.4537 | -35.5297 | -34.6937
24.4598 | -35.5301 | -34.699
24.4659 | -35.5305 | -34.7047
24.472 | -35.5309 | -34.7099
24.4781 | -35.5313 | -34.7153
24.4842 | -35.5317 | -34.7205
24.4903 | -35.532 | -34.7262
24.4964 | -35.5324 | -34.7321
24.5025 | -35.5329 | -34.7377
24.5086 | -35.5333 | -34.7433
24.5147 | -35.5337 | -34.7488
24.5208 | -35.5342 | -34.7542
24.5269 | -35.5346 | -34.7601
24.533 | -35.535 | -34.7658
24.5391 | -35.5355 | -34.7713
24.5452 | -35.5359 | -34.777
24.5513 | -35.5363 | -34.7824
24.5574 | -35.5367 | -34.7878
24.5635 | -35.5371 | -34.7936
24.5696 | -35.5375 | -34.7991
24.5758 | -35.5379 | -34.8045
24.5819 | -35.5383 | -34.8101
24.588 | -35.5386 | -34.8151
24.5941 | -35.539 | -34.8204
24.6002 | -35.5393 | -34.8258
24.6063 | -35.5398 | -34.8313
24.6124 | -35.5402 | -34.8369
24.6185 | -35.5406 | -34.8427
24.6246 | -35.541 | -34.8479
24.6307 | -35.5413 | -34.8529
24.6368 | -35.5417 | -34.8588
24.6429 | -35.5421 | -34.8643
24.649 | -35.5425 | -34.8697
24.6551 | -35.5429 | -34.8752
24.6612 | -35.5433 | -34.8806
24.6673 | -35.5436 | -34.8859
24.6734 | -35.544 | -34.8917
24.6795 | -35.5443 | -34.8979
24.6856 | -35.5447 | -34.9032
24.6917 | -35.545 | -34.9084
24.6978 | -35.5454 | -34.914
24.7039 | -35.5458 | -34.9197
24.71 | -35.5461 | -34.9255
24.7162 | -35.5465 | -34.931
24.7223 | -35.5469 | -34.9363
24.7284 | -35.5472 | -34.9422
24.7345 | -35.5476 | -34.9475
24.7406 | -35.5479 | -34.953
24.7467 | -35.5483 | -34.9586
24.7528 | -35.5487 | -34.9641
24.7589 | -35.5491 | -34.9691
24.765 | -35.5494 | -34.9745
24.7711 | -35.5498 | -34.9802
24.7772 | -35.5502 | -34.9858
24.7833 | -35.5506 | -34.9913
24.7894 | -35.5509 | -34.9967
24.7955 | -35.5513 | -35.002
24.8016 | -35.5518 | -35.008
24.8077 | -35.5522 | -35.0132
24.8138 | -35.5526 | -35.0186
24.8199 | -35.553 | -35.0239
24.826 | -35.5534 | -35.0295
24.8321 | -35.5537 | -35.0351
24.8382 | -35.5541 | -35.0405
24.8443 | -35.5545 | -35.0457
24.8504 | -35.5549 | -35.0511
24.8565 | -35.5553 | -35.0569
24.8627 | -35.5556 | -35.0623
24.8688 | -35.556 | -35.0671
24.8749 | -35.5564 | -35.0724
24.881 | -35.5568 | -35.0779
24.8871 | -35.5572 | -35.0833
24.8932 | -35.5576 | -35.0889
24.8993 | -35.558 | -35.0943
24.9054 | -35.5584 | -35.0995
24.9115 | -35.5588 | -35.1048
24.9176 | -35.5592 | -35.1101
24.9237 | -35.5596 | -35.1156
24.9298 | -35.5599 | -35.1211
24.9359 | -35.5603 | -35.1267
24.942 | -35.5607 | -35.1316
24.9481 | -35.561 | -35.1369
24.9542 | -35.5614 | -35.1424
24.9603 | -35.5617 | -35.1479
24.9664 | -35.5621 | -35.1533
24.9725 | -35.5626 | -35.1588
24.9786 | -35.5629 | -35.1639
24.9847 | -35.5633 | -35.1694
24.9908 | -35.5637 | -35.175
24.9969 | -35.564 | -35.1807
25.0031 | -35.5644 | -35.1862
25.0092 | -35.5648 | -35.1915
25.0153 | -35.5652 | -35.1967
25.0214 | -35.5656 | -35.202
25.0275 | -35.566 | -35.2075
25.0336 | -35.5664 | -35.213
25.0397 | -35.5668 | -35.2187
25.0458 | -35.5672 | -35.2239
25.0519 | -35.5676 | -35.2292
25.058 | -35.568 | -35.2345
25.0641 | -35.5683 | -35.24
25.0702 | -35.5687 | -35.2452
25.0763 | -35.5691 | -35.2506
25.0824 | -35.5695 | -35.2556
25.0885 | -35.5698 | -35.2609
25.0946 | -35.5702 | -35.2664
25.1007 | -35.5706 | -35.2719
25.1068 | -35.571 | -35.277
25.1129 | -35.5714 | -35.2826
25.119 | -35.5718 | -35.2879
25.1251 | -35.5722 | -35.2935
25.1312 | -35.5725 | -35.2995
25.1373 | -35.5729 | -35.305
25.1435 | -35.5733 | -35.3107
25.1496 | -35.5737 | -35.3154
25.1557 | -35.574 | -35.3208
25.1618 | -35.5744 | -35.3263
25.1679 | -35.5747 | -35.3318
25.174 | -35.5751 | -35.3373
25.1801 | -35.5754 | -35.3425
25.1862 | -35.5758 | -35.3481
25.1923 | -35.5761 | -35.3533
25.1984 | -35.5766 | -35.3588
25.2045 | -35.577 | -35.3643
25.2106 | -35.5773 | -35.3698
25.2167 | -35.5777 | -35.3752
25.2228 | -35.5779 | -35.3807
25.2289 | -35.5783 | -35.386
25.235 | -35.5787 | -35.3915
25.2411 | -35.5791 | -35.3969
25.2472 | -35.5795 | -35.4022
25.2533 | -35.5798 | -35.4079
25.2594 | -35.5803 | -35.4133
25.2655 | -35.5807 | -35.4188
25.2716 | -35.581 | -35.4242
25.2777 | -35.5815 | -35.4297
25.2838 | -35.5818 | -35.435
25.29 | -35.5822 | -35.4402
25.2961 | -35.5826 | -35.4457
25.3022 | -35.5831 | -35.4512
25.3083 | -35.5835 | -35.4567
25.3144 | -35.5839 | -35.4621
25.3205 | -35.5842 | -35.4675
25.3266 | -35.5846 | -35.4728
25.3327 | -35.5849 | -35.4783
25.3388 | -35.5853 | -35.4837
25.3449 | -35.5857 | -35.489
25.351 | -35.586 | -35.4942
25.3571 | -35.5864 | -35.5
25.3632 | -35.5868 | -35.5053
25.3693 | -35.5872 | -35.5111
25.3754 | -35.5876 | -35.5166
25.3815 | -35.5879 | -35.5221
25.3876 | -35.5883 | -35.5274
25.3937 | -35.5887 | -35.5328
25.3998 | -35.5891 | -35.5382
25.4059 | -35.5894 | -35.5438
25.412 | -35.5898 | -35.5489
25.4181 | -35.5902 | -35.5543
25.4242 | -35.5905 | -35.5597
25.4304 | -35.5908 | -35.5651
25.4365 | -35.5912 | -35.5703
25.4426 | -35.5916 | -35.5755
25.4487 | -35.592 | -35.581
25.4548 | -35.5923 | -35.5865
25.4609 | -35.5927 | -35.5921
25.467 | -35.5931 | -35.5974
25.4731 | -35.5935 | -35.6028
25.4792 | -35.5938 | -35.6082
25.4853 | -35.5942 | -35.6137
25.4914 | -35.5946 | -35.619
25.4975 | -35.595 | -35.6245
25.5036 | -35.5954 | -35.6301
25.5097 | -35.5958 | -35.6355
25.5158 | -35.5962 | -35.6409
25.5219 | -35.5966 | -35.6464
25.528 | -35.5969 | -35.6518
25.5341 | -35.5973 | -35.6574
25.5402 | -35.5977 | -35.663
25.5463 | -35.5981 | -35.668
25.5524 | -35.5985 | -35.6735
25.5585 | -35.5988 | -35.6785
25.5646 | -35.5992 | -35.6839
25.5707 | -35.5995 | -35.6892
25.5769 | -35.5999 | -35.6948
25.583 | -35.6003 | -35.6999
25.5891 | -35.6007 | -35.705
25.5952 | -35.6011 | -35.7108
25.6013 | -35.6015 | -35.7164
25.6074 | -35.6019 | -35.7217
25.6135 | -35.6023 | -35.7269
25.6196 | -35.6027 | -35.7325
25.6257 | -35.603 | -35.7378
25.6318 | -35.6034 | -35.7429
25.6379 | -35.6038 | -35.7482
25.644 | -35.6042 | -35.7535
25.6501 | -35.6045 | -35.7592
25.6562 | -35.6049 | -35.7645
25.6623 | -35.6052 | -35.7696
25.6684 | -35.6056 | -35.7751
25.6745 | -35.606 | -35.7805
25.6806 | -35.6065 | -35.7855
25.6867 | -35.6068 | -35.7909
25.6928 | -35.6072 | -35.7962
25.6989 | -35.6076 | -35.8017
25.705 | -35.608 | -35.8069
25.7111 | -35.6083 | -35.8124
25.7173 | -35.6087 | -35.8179
25.7234 | -35.6091 | -35.8233
25.7295 | -35.6095 | -35.8286
25.7356 | -35.6099 | -35.8343
25.7417 | -35.6102 | -35.8392
25.7478 | -35.6106 | -35.8444
25.7539 | -35.6109 | -35.8497
25.76 | -35.6112 | -35.8549
25.7661 | -35.6116 | -35.8602
25.7722 | -35.612 | -35.8653
25.7783 | -35.6124 | -35.8709
25.7844 | -35.6127 | -35.8765
25.7905 | -35.613 | -35.8818
25.7966 | -35.6134 | -35.8869
25.8027 | -35.6137 | -35.8923
25.8088 | -35.6141 | -35.8976
25.8149 | -35.6145 | -35.9028
25.821 | -35.6148 | -35.9078
25.8271 | -35.6153 | -35.9134
25.8332 | -35.6157 | -35.9188
25.8393 | -35.6161 | -35.9241
25.8454 | -35.6164 | -35.9295
25.8515 | -35.6167 | -35.9351
25.8576 | -35.617 | -35.9404
25.8638 | -35.6174 | -35.9459
25.8699 | -35.6178 | -35.9513
25.876 | -35.6182 | -35.9565
25.8821 | -35.6185 | -35.9617
25.8882 | -35.6189 | -35.967
25.8943 | -35.6193 | -35.9726
25.9004 | -35.6196 | -35.978
25.9065 | -35.6199 | -35.9834
25.9126 | -35.6203 | -35.9881
25.9187 | -35.6207 | -35.9934
25.9248 | -35.621 | -35.9984
25.9309 | -35.6215 | -36.004
25.937 | -35.6218 | -36.0094
25.9431 | -35.6222 | -36.0146
25.9492 | -35.6225 | -36.0199
25.9553 | -35.6229 | -36.0252
25.9614 | -35.6232 | -36.0302
25.9675 | -35.6236 | -36.0351
25.9736 | -35.624 | -36.0405
25.9797 | -35.6243 | -36.0458
25.9858 | -35.6247 | -36.051
25.9919 | -35.625 | -36.0563
25.998 | -35.6254 | -36.0618
26.0042 | -35.6258 | -36.0671
26.0103 | -35.6262 | -36.0724
26.0164 | -35.6265 | -36.0779
26.0225 | -35.6269 | -36.0831
26.0286 | -35.6272 | -36.0885
26.0347 | -35.6276 | -36.094
26.0408 | -35.6279 | -36.0993
26.0469 | -35.6283 | -36.1046
26.053 | -35.6286 | -36.1101
26.0591 | -35.629 | -36.1154
26.0652 | -35.6293 | -36.1206
26.0713 | -35.6297 | -36.126
26.0774 | -35.6301 | -36.1311
26.0835 | -35.6305 | -36.1363
26.0896 | -35.6309 | -36.1417
26.0957 | -35.6312 | -36.147
26.1018 | -35.6316 | -36.1521
26.1079 | -35.632 | -36.1574
26.114 | -35.6323 | -36.1627
26.1201 | -35.6327 | -36.168
26.1262 | -35.633 | -36.1731
26.1323 | -35.6334 | -36.1783
26.1384 | -35.6338 | -36.1831
26.1445 | -35.6342 | -36.1885
26.1507 | -35.6345 | -36.1941
26.1568 | -35.6349 | -36.1991
26.1629 | -35.6353 | -36.2045
26.169 | -35.6357 | -36.2097
26.1751 | -35.636 | -36.2146
26.1812 | -35.6364 | -36.22
26.1873 | -35.6367 | -36.2254
26.1934 | -35.637 | -36.2305
26.1995 | -35.6374 | -36.2359
26.2056 | -35.6378 | -36.2408
26.2117 | -35.6382 | -36.2459
26.2178 | -35.6385 | -36.251
26.2239 | -35.6389 | -36.2563
26.23 | -35.6393 | -36.2616
26.2361 | -35.6396 | -36.2671
26.2422 | -35.64 | -36.2722
26.2483 | -35.6403 | -36.2775
26.2544 | -35.6407 | -36.2827
26.2605 | -35.641 | -36.2881
26.2666 | -35.6414 | -36.2936
26.2727 | -35.6418 | -36.299
26.2788 | -35.6421 | -36.3048
26.2849 | -35.6425 | -36.3101
26.2911 | -35.6429 | -36.3151
26.2972 | -35.6434 | -36.3204
26.3033 | -35.6438 | -36.326
26.3094 | -35.6441 | -36.3309
26.3155 | -35.6444 | -36.3364
26.3216 | -35.6448 | -36.342
26.3277 | -35.6452 | -36.3469
26.3338 | -35.6455 | -36.3518
26.3399 | -35.6459 | -36.3569
26.346 | -35.6463 | -36.3621
26.3521 | -35.6466 | -36.3675
26.3582 | -35.647 | -36.3726
26.3643 | -35.6474 | -36.3778
26.3704 | -35.6477 | -36.3832
26.3765 | -35.6481 | -36.3884
26.3826 | -35.6484 | -36.3935
26.3887 | -35.6489 | -36.3991
26.3948 | -35.6492 | -36.4044
26.4009 | -35.6496 | -36.4096
26.407 | -35.65 | -36.4148
26.4131 | -35.6503 | -36.4202
26.4192 | -35.6507 | -36.425
26.4253 | -35.6511 | -36.4307
26.4314 | -35.6515 | -36.4359
26.4376 | -35.6519 | -36.4411
26.4437 | -35.6522 | -36.4463
26.4498 | -35.6526 | -36.4515
26.4559 | -35.6529 | -36.4566
26.462 | -35.6533 | -36.4619
26.4681 | -35.6537 | -36.4674
26.4742 | -35.654 | -36.4726
26.4803 | -35.6544 | -36.4778
26.4864 | -35.6547 | -36.4829
26.4925 | -35.6551 | -36.4879
26.4986 | -35.6555 | -36.4934
26.5047 | -35.6559 | -36.4986
26.5108 | -35.6563 | -36.504
26.5169 | -35.6566 | -36.5091
26.523 | -35.657 | -36.5142
26.5291 | -35.6574 | -36.5191
26.5352 | -35.6578 | -36.5238
26.5413 | -35.6581 | -36.5296
26.5474 | -35.6584 | -36.5347
26.5535 | -35.6588 | -36.5402
26.5596 | -35.6591 | -36.5455
26.5657 | -35.6595 | -36.5509
26.5718 | -35.6599 | -36.5564
26.578 | -35.6603 | -36.5618
26.5841 | -35.6607 | -36.5669
26.5902 | -35.661 | -36.5719
26.5963 | -35.6614 | -36.5771
26.6024 | -35.6617 | -36.5827
26.6085 | -35.6621 | -36.5879
26.6146 | -35.6624 | -36.5934
26.6207 | -35.6628 | -36.5987
26.6268 | -35.6632 | -36.6036
26.6329 | -35.6635 | -36.6086
26.639 | -35.6639 | -36.6137
26.6451 | -35.6643 | -36.6191
26.6512 | -35.6646 | -36.6242
26.6573 | -35.6649 | -36.6296
26.6634 | -35.6653 | -36.6341
26.6695 | -35.6656 | -36.6397
26.6756 | -35.6661 | -36.6452
26.6817 | -35.6664 | -36.6502
26.6878 | -35.6667 | -36.6553
26.6939 | -35.6671 | -36.6606
26.7 | -35.6674 | -36.6653
26.7061 | -35.6678 | -36.6702
26.7122 | -35.6681 | -36.6757
26.7183 | -35.6685 | -36.681
26.7245 | -35.6688 | -36.6859
26.7306 | -35.6691 | -36.6908
26.7367 | -35.6695 | -36.6957
26.7428 | -35.6698 | -36.701
26.7489 | -35.6702 | -36.7064
26.755 | -35.6706 | -36.7117
26.7611 | -35.6709 | -36.7172
26.7672 | -35.6712 | -36.7221
26.7733 | -35.6716 | -36.7276
26.7794 | -35.672 | -36.7329
26.7855 | -35.6724 | -36.7382
26.7916 | -35.6727 | -36.7431
26.7977 | -35.6731 | -36.7484
26.8038 | -35.6735 | -36.7533
26.8099 | -35.6738 | -36.7581
26.816 | -35.6742 | -36.7635
26.8221 | -35.6746 | -36.7691
26.8282 | -35.675 | -36.7739
26.8343 | -35.6754 | -36.7791
26.8404 | -35.6757 | -36.7841
26.8465 | -35.676 | -36.7894
26.8526 | -35.6764 | -36.7949
26.8587 | -35.6768 | -36.8
26.8649 | -35.6771 | -36.8053
26.871 | -35.6775 | -36.8102
26.8771 | -35.6779 | -36.8152
26.8832 | -35.6783 | -36.8204
26.8893 | -35.6786 | -36.8255
26.8954 | -35.679 | -36.8306
26.9015 | -35.6794 | -36.836
26.9076 | -35.6797 | -36.8418
26.9137 | -35.6801 | -36.8468
26.9198 | -35.6804 | -36.8522
26.9259 | -35.6808 | -36.8575
26.932 | -35.6811 | -36.8629
26.9381 | -35.6814 | -36.8682
26.9442 | -35.6818 | -36.8734
26.9503 | -35.6821 | -36.8784
26.9564 | -35.6825 | -36.8834
26.9625 | -35.6829 | -36.8886
26.9686 | -35.6832 | -36.8934
26.9747 | -35.6836 | -36.8982
26.9808 | -35.6839 | -36.9029
26.9869 | -35.6842 | -36.908
26.993 | -35.6846 | -36.9135
26.9991 | -35.6849 | -36.9184
27.0052 | -35.6853 | -36.9236
27.0114 | -35.6857 | -36.929
27.0175 | -35.686 | -36.9344
27.0236 | -35.6863 | -36.9396
27.0297 | -35.6866 | -36.9447
27.0358 | -35.687 | -36.95
27.0419 | -35.6874 | -36.955
27.048 | -35.6877 | -36.9602
27.0541 | -35.688 | -36.9653
27.0602 | -35.6883 | -36.9705
27.0663 | -35.6887 | -36.9757
27.0724 | -35.689 | -36.9809
27.0785 | -35.6894 | -36.9863
27.0846 | -35.6897 | -36.991
27.0907 | -35.6901 | -36.9958
27.0968 | -35.6904 | -37.0008
27.1029 | -35.6908 | -37.006
27.109 | -35.6911 | -37.0108
27.1151 | -35.6915 | -37.0161
27.1212 | -35.6918 | -37.0212
27.1273 | -35.6921 | -37.0263
27.1334 | -35.6925 | -37.0311
27.1395 | -35.6929 | -37.0362
27.1456 | -35.6933 | -37.0415
27.1518 | -35.6937 | -37.047
27.1579 | -35.694 | -37.0519
27.164 | -35.6943 | -37.0571
27.1701 | -35.6947 | -37.0619
27.1762 | -35.695 | -37.0669
27.1823 | -35.6954 | -37.0716
27.1884 | -35.6958 | -37.0766
27.1945 | -35.6962 | -37.0814
27.2006 | -35.6965 | -37.0867
27.2067 | -35.6969 | -37.092
27.2128 | -35.6972 | -37.0972
27.2189 | -35.6975 | -37.1026
27.225 | -35.6979 | -37.108
27.2311 | -35.6983 | -37.113
27.2372 | -35.6986 | -37.1181
27.2433 | -35.6989 | -37.1231
27.2494 | -35.6993 | -37.1283
27.2555 | -35.6996 | -37.1336
27.2616 | -35.7 | -37.1387
27.2677 | -35.7003 | -37.1434
27.2738 | -35.7007 | -37.1485
27.2799 | -35.701 | -37.1538
27.286 | -35.7013 | -37.1588
27.2921 | -35.7016 | -37.164
27.2983 | -35.702 | -37.1687
27.3044 | -35.7024 | -37.1739
27.3105 | -35.7028 | -37.1791
27.3166 | -35.7032 | -37.1844
27.3227 | -35.7035 | -37.1894
27.3288 | -35.7038 | -37.1948
27.3349 | -35.7041 | -37.1997
27.341 | -35.7044 | -37.2045
27.3471 | -35.7048 | -37.2094
27.3532 | -35.7051 | -37.2143
27.3593 | -35.7054 | -37.2196
27.3654 | -35.7058 | -37.2251
27.3715 | -35.7061 | -37.2301
27.3776 | -35.7064 | -37.2357
27.3837 | -35.7068 | -37.2413
27.3898 | -35.7071 | -37.246
27.3959 | -35.7074 | -37.251
27.402 | -35.7077 | -37.2557
27.4081 | -35.708 | -37.2605
27.4142 | -35.7083 | -37.2659
27.4203 | -35.7087 | -37.2706
27.4264 | -35.709 | -37.276
27.4325 | -35.7094 | -37.2812
27.4387 | -35.7097 | -37.2861
27.4448 | -35.71 | -37.2915
27.4509 | -35.7103 | -37.2969
27.457 | -35.7106 | -37.3019
27.4631 | -35.711 | -37.3068
27.4692 | -35.7113 | -37.312
27.4753 | -35.7117 | -37.3167
27.4814 | -35.7121 | -37.3214
27.4875 | -35.7124 | -37.3264
27.4936 | -35.7128 | -37.3317
27.4997 | -35.7131 | -37.3366
27.5058 | -35.7134 | -37.3413
27.5119 | -35.7137 | -37.3462
27.518 | -35.714 | -37.3514
27.5241 | -35.7143 | -37.3562
27.5302 | -35.7147 | -37.3613
27.5363 | -35.715 | -37.3664
27.5424 | -35.7153 | -37.3712
27.5485 | -35.7157 | -37.3761
27.5546 | -35.716 | -37.3808
27.5607 | -35.7163 | -37.3856
27.5668 | -35.7167 | -37.3903
27.5729 | -35.7171 | -37.3953
27.5791 | -35.7174 | -37.4004
27.5852 | -35.7178 | -37.4052
27.5913 | -35.7181 | -37.41
27.5974 | -35.7184 | -37.4152
27.6035 | -35.7187 | -37.4199
27.6096 | -35.7191 | -37.4247
27.6157 | -35.7194 | -37.4298
27.6218 | -35.7198 | -37.4348
27.6279 | -35.7202 | -37.4395
27.634 | -35.7205 | -37.4448
27.6401 | -35.7209 | -37.4497
27.6462 | -35.7213 | -37.4548
27.6523 | -35.7216 | -37.4595
27.6584 | -35.722 | -37.4641
27.6645 | -35.7224 | -37.4693
27.6706 | -35.7227 | -37.4742
27.6767 | -35.723 | -37.4792
27.6828 | -35.7233 | -37.4848
27.6889 | -35.7237 | -37.4897
27.695 | -35.724 | -37.4947
27.7011 | -35.7244 | -37.4995
27.7072 | -35.7247 | -37.5049
27.7133 | -35.725 | -37.5099
27.7194 | -35.7254 | -37.5154
27.7256 | -35.7257 | -37.5203
27.7317 | -35.726 | -37.5251
27.7378 | -35.7264 | -37.5302
27.7439 | -35.7267 | -37.5355
27.75 | -35.7271 | -37.5406
27.7561 | -35.7274 | -37.5453
27.7622 | -35.7278 | -37.5506
27.7683 | -35.7281 | -37.5559
27.7744 | -35.7285 | -37.5609
27.7805 | -35.7288 | -37.5661
27.7866 | -35.729 | -37.5713
27.7927 | -35.7293 | -37.5755
27.7988 | -35.7297 | -37.5805
27.8049 | -35.73 | -37.5856
27.811 | -35.7304 | -37.5903
27.8171 | -35.7307 | -37.5956
27.8232 | -35.7311 | -37.6004
27.8293 | -35.7314 | -37.6053
27.8354 | -35.7317 | -37.6106
27.8415 | -35.732 | -37.6154
27.8476 | -35.7323 | -37.6203
27.8537 | -35.7327 | -37.6253
27.8598 | -35.733 | -37.6298
27.866 | -35.7333 | -37.6355
27.8721 | -35.7336 | -37.6404
27.8782 | -35.7339 | -37.6454
27.8843 | -35.7342 | -37.6505
27.8904 | -35.7345 | -37.656
27.8965 | -35.7348 | -37.6608
27.9026 | -35.7352 | -37.666
27.9087 | -35.7356 | -37.6709
27.9148 | -35.736 | -37.6756
27.9209 | -35.7363 | -37.6806
27.927 | -35.7366 | -37.6853
27.9331 | -35.737 | -37.6901
27.9392 | -35.7373 | -37.6957
27.9453 | -35.7376 | -37.7003
27.9514 | -35.738 | -37.7054
27.9575 | -35.7383 | -37.7105
27.9636 | -35.7386 | -37.7151
27.9697 | -35.739 | -37.7192
27.9758 | -35.7393 | -37.724
27.9819 | -35.7396 | -37.729
27.988 | -35.74 | -37.7337
27.9941 | -35.7403 | -37.7385
28.0002 | -35.7406 | -37.7426
28.0063 | -35.7409 | -37.7476
28.0125 | -35.7411 | -37.7523
28.0186 | -35.7415 | -37.7574
28.0247 | -35.7418 | -37.7623
28.0308 | -35.7421 | -37.7668
28.0369 | -35.7424 | -37.7713
28.043 | -35.7426 | -37.7766
28.0491 | -35.743 | -37.7816
28.0552 | -35.7434 | -37.7861
28.0613 | -35.7437 | -37.7904
28.0674 | -35.744 | -37.7956
28.0735 | -35.7443 | -37.7998

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.3467 | -42.2206
21.4908 | -35.6129 | -34.9358
16.4948 | -34.9731 | -27.6248
12.6602 | -34.6067 | -21.3915
9.7171 | -34.3865 | -16.5127
7.4582 | -34.2325 | -12.549
5.7244 | -34.1219 | -9.3389
4.3936 | -34.0567 | -6.748
3.3722 | -34.0412 | -4.7047
2.5883 | -34.0422 | -3.1539
1.9866 | -34.0385 | -1.9158
1.5248 | -34.0306 | -0.91848
1.1703 | -34.0376 | -0.27034
1 | -34.0306 | 0.05258
0.89822 | -34.0329 | 0.32536
0.68941 | -34.0152 | 0.90706
0.52919 | -34.0243 | 1.3596
0.40612 | -34.0081 | 1.9659
0.31171 | -33.9904 | 2.4214
0.23928 | -34.0069 | 3.1498
0.18358 | -34.0134 | 3.5689
0.14091 | -34.0357 | 4.4468
0.10819 | -34.0489 | 5.6161
0.083017 | -34.0772 | 7.1013
0.063705 | -34.11 | 8.9767
0.048923 | -34.1713 | 11.4799
0.037479 | -34.2878 | 14.6705
0.028801 | -34.4805 | 18.7008
0.022078 | -34.7778 | 24.0765
0.017023 | -35.2593 | 30.5015
0.013018 | -36.064 | 38.2117
0.010014 | -37.1872 | 46.1403

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003775 | -1.2321
0.007549 | -25.2301
0.011324 | -33.8343
0.015099 | -40.2453
0.018873 | -44.1617
0.022648 | -45.3572
0.026423 | -47.3448
0.030197 | -48.398
0.033972 | -48.1051
0.037747 | -48.4624
0.041521 | -50.1564
0.045296 | -53.0752
0.04907 | -54.7145
0.052845 | -55.4103
0.05662 | -56.2495
0.060394 | -56.561
0.064169 | -57.4286
0.067944 | -58.0838
0.071718 | -58.5202
0.075493 | -59.1996
0.079268 | -59.6131
0.083042 | -59.927
0.086817 | -59.9155
0.090592 | -61.4041
0.094366 | -61.7249
0.098141 | -60.8823
0.10192 | -62.0879
0.10569 | -63.7211
0.10947 | -64.4655
0.11324 | -65.1634
0.11701 | -64.6135
0.12079 | -65.1555
0.12456 | -66.9014
0.12834 | -66.3588
0.13211 | -66.7437
0.13589 | -67.0649
0.13966 | -67.8877
0.14344 | -67.9691
0.14721 | -67.4443
0.15099 | -68.6551
0.15476 | -67.7752
0.15854 | -67.8488
0.16231 | -68.5874
0.16609 | -68.6023
0.16986 | -67.8699
0.17363 | -68.8561
0.17741 | -69.6076
0.18118 | -69.364
0.18496 | -69.7432
0.18873 | -69.2926
0.19251 | -70.3382
0.19628 | -70.7204
0.20006 | -71.0305
0.20383 | -71.4886
0.20761 | -70.8304
0.21138 | -71.6207
0.21516 | -71.4801
0.21893 | -72.1877
0.2227 | -72.3762
0.22648 | -72.4376
0.23025 | -72.5465
0.23403 | -72.5397
0.2378 | -72.805
0.24158 | -72.4109
0.24535 | -73.6344
0.24913 | -73.7953
0.2529 | -73.5686
0.25668 | -73.9595
0.26045 | -74.104
0.26423 | -74.708
0.268 | -74.4709
0.27177 | -74.0659
0.27555 | -74.7188
0.27932 | -75.0103
0.2831 | -74.5789
0.28687 | -74.7981
0.29065 | -74.6917
0.29442 | -74.5809
0.2982 | -74.7153
0.30197 | -75.5213
0.30575 | -76.3523
0.30952 | -76.9017
0.3133 | -76.3428
0.31707 | -76.6988
0.32085 | -77.3599
0.32462 | -77.0353
0.32839 | -76.9727
0.33217 | -78.5437
0.33594 | -77.6913
0.33972 | -76.6364
0.34349 | -77.1581
0.34727 | -77.3351
0.35104 | -77.1284
0.35482 | -77.6024
0.35859 | -77.2903
0.36237 | -77.3825
0.36614 | -78.4535
0.36992 | -78.2546
0.37369 | -78.3429
0.37746 | -79.1036
0.38124 | -78.7203
0.38501 | -78.3631
0.38879 | -77.4078
0.39256 | -77.9968
0.39634 | -77.5188
0.40011 | -77.4934
0.40389 | -78.5335
0.40766 | -78.7623
0.41144 | -78.2297
0.41521 | -79.385
0.41899 | -79.5215
0.42276 | -78.8027
0.42654 | -79.9152
0.43031 | -79.8449
0.43408 | -79.6546
0.43786 | -78.9321
0.44163 | -78.4591
0.44541 | -78.5042
0.44918 | -79.7366
0.45296 | -80.6157
0.45673 | -81.2464
0.46051 | -80.6177
0.46428 | -78.8115
0.46806 | -78.4888
0.47183 | -79.0068
0.47561 | -79.1955
0.47938 | -78.6102
0.48316 | -80.3829
0.48693 | -81.7675
0.49071 | -81.3806
0.49448 | -81.5094
0.49825 | -81.1958
0.50203 | -80.6717
0.5058 | -80.7755
0.50958 | -81.0872
0.51335 | -81.6944
0.51713 | -82.0289
0.5209 | -82.2968
0.52468 | -81.7294
0.52845 | -81.7746
0.53223 | -82.7579
0.536 | -83.4782
0.53978 | -81.8115
0.54355 | -80.9735
0.54732 | -80.8311
0.5511 | -80.8978
0.55487 | -80.7473
0.55865 | -81.1956
0.56242 | -81.9806
0.5662 | -82.3742
0.56997 | -82.1912
0.57375 | -81.986
0.57752 | -82.3467
0.5813 | -82.7604
0.58507 | -82.2821
0.58885 | -81.8206
0.59262 | -82.071
0.5964 | -82.6862
0.60017 | -82.9423
0.60394 | -82.6606
0.60772 | -82.3441
0.61149 | -82.2691
0.61527 | -82.4278
0.61904 | -82.744
0.62282 | -82.8978
0.62659 | -82.8877
0.63037 | -82.662
0.63414 | -82.2426
0.63792 | -82.2211
0.64169 | -82.2293
0.64547 | -82.4306
0.64924 | -83.2396
0.65302 | -83.5995
0.65679 | -83.4988
0.66056 | -83.3084
0.66434 | -83.1183
0.66811 | -83.151
0.67189 | -83.8025
0.67566 | -84.5374
0.67944 | -84.5647
0.68321 | -84.5011
0.68699 | -84.4196
0.69076 | -83.9486
0.69454 | -83.597
0.69831 | -83.8706
0.70208 | -84.685
0.70586 | -85.1469
0.70964 | -84.8616
0.71341 | -84.6397
0.71718 | -84.4606
0.72096 | -84.4052
0.72473 | -84.727
0.72851 | -84.6641
0.73228 | -84.0342
0.73606 | -83.5853
0.73983 | -83.3276
0.74361 | -83.3695
0.74738 | -83.7528
0.75116 | -83.9621
0.75493 | -84.094
0.7587 | -84.4914
0.76248 | -84.8048
0.76626 | -84.8565
0.77003 | -84.9221
0.7738 | -85.0819
0.77758 | -85.5748
0.78135 | -85.6529
0.78513 | -85.3354
0.7889 | -84.9815
0.79268 | -84.8869
0.79645 | -85.17
0.80023 | -85.6941
0.804 | -86.0551
0.80778 | -85.7737
0.81155 | -85.1711
0.81532 | -85.2457
0.8191 | -85.697
0.82287 | -85.5869
0.82665 | -85.7857
0.83042 | -86.0659
0.8342 | -86.054
0.83797 | -85.9366
0.84175 | -85.7074
0.84552 | -85.3499
0.8493 | -85.3386
0.85307 | -85.5754
0.85685 | -85.5236
0.86062 | -85.4281
0.8644 | -85.5479
0.86817 | -85.695
0.87194 | -85.851
0.87572 | -86.1571
0.87949 | -86.7207
0.88327 | -87.2229
0.88704 | -86.7258
0.89082 | -86.1212
0.89459 | -86.1498
0.89837 | -86.4857
0.90214 | -87.1205
0.90592 | -87.6864
0.90969 | -87.5132
0.91347 | -87.3037
0.91724 | -87.4135
0.92102 | -87.1479
0.92479 | -86.1766
0.92856 | -85.5669
0.93234 | -85.6839
0.93611 | -86.2269
0.93989 | -87.0281
0.94366 | -87.2165
0.94744 | -87.1372
0.95121 | -87.2794
0.95499 | -87.1757
0.95876 | -86.9748
0.96254 | -86.7262
0.96631 | -86.3739
0.97009 | -86.5924
0.97386 | -87.0408
0.97764 | -86.9793
0.98141 | -86.9042
0.98518 | -86.8314
0.98896 | -86.8515
0.99273 | -87.2396
0.99651 | -87.6049
1.0003 | -87.5617
1.0041 | -87.7173
1.0078 | -87.4737
1.0116 | -86.8272
1.0154 | -86.7229
1.0192 | -87.1693
1.0229 | -87.302
1.0267 | -87.3802
1.0305 | -87.5153
1.0343 | -87.4478
1.038 | -87.6421
1.0418 | -87.5295
1.0456 | -87.2928
1.0494 | -87.5887
1.0531 | -88.1978
1.0569 | -88.6253
1.0607 | -88.7472
1.0645 | -88.4882
1.0682 | -88.5139
1.072 | -88.7909
1.0758 | -88.5044
1.0796 | -88.195
1.0833 | -88.0357
1.0871 | -87.9884
1.0909 | -88.131
1.0946 | -88.3957
1.0984 | -88.654
1.1022 | -88.7509
1.106 | -88.5797
1.1097 | -88.3748
1.1135 | -88.3345
1.1173 | -88.3174
1.1211 | -88.1029
1.1248 | -87.8693
1.1286 | -87.9274
1.1324 | -88.3066
1.1362 | -88.8704
1.1399 | -89.4479
1.1437 | -89.8265
1.1475 | -89.817
1.1513 | -89.3878
1.155 | -88.803
1.1588 | -88.4315
1.1626 | -88.4279
1.1664 | -88.5756
1.1701 | -88.6181
1.1739 | -88.5505
1.1777 | -88.4554
1.1815 | -88.4203
1.1852 | -88.4575
1.189 | -88.5234
1.1928 | -88.6419
1.1966 | -88.77
1.2003 | -88.8174
1.2041 | -88.8346
1.2079 | -88.8374
1.2117 | -88.7687
1.2154 | -88.6387
1.2192 | -88.538
1.223 | -88.5628
1.2268 | -88.6749
1.2305 | -88.8136
1.2343 | -89.0176
1.2381 | -89.2397
1.2419 | -89.4167
1.2456 | -89.5946
1.2494 | -89.8347
1.2532 | -90.016
1.257 | -89.9083
1.2607 | -89.4674
1.2645 | -88.9769
1.2683 | -88.8111
1.2721 | -88.9392
1.2758 | -89.0259
1.2796 | -89.103
1.2834 | -89.3982
1.2872 | -89.8009
1.2909 | -89.9835
1.2947 | -89.8748
1.2985 | -89.7516
1.3023 | -89.8077
1.306 | -89.9583
1.3098 | -90.0962
1.3136 | -90.1322
1.3174 | -90.0739
1.3211 | -90.1443
1.3249 | -90.3977
1.3287 | -90.5972
1.3325 | -90.6351
1.3362 | -90.5115
1.34 | -90.2521
1.3438 | -89.9695
1.3476 | -89.8141
1.3513 | -89.8705
1.3551 | -90.1073
1.3589 | -90.3734
1.3626 | -90.5188
1.3664 | -90.4662
1.3702 | -90.2955
1.374 | -90.2191
1.3777 | -90.3286
1.3815 | -90.4523
1.3853 | -90.5338
1.3891 | -90.7356
1.3928 | -90.9774
1.3966 | -91.0047
1.4004 | -90.844
1.4042 | -90.7084
1.4079 | -90.6057
1.4117 | -90.4854
1.4155 | -90.4593
1.4193 | -90.4909
1.423 | -90.4581
1.4268 | -90.4601
1.4306 | -90.5767
1.4344 | -90.739
1.4381 | -90.7782
1.4419 | -90.6324
1.4457 | -90.5677
1.4495 | -90.6718
1.4532 | -90.5978
1.457 | -90.2395
1.4608 | -89.9733
1.4646 | -90.0514
1.4683 | -90.3448
1.4721 | -90.5705
1.4759 | -90.5649
1.4797 | -90.333
1.4834 | -90.0883
1.4872 | -90.011
1.491 | -90.0435
1.4948 | -90.0936
1.4985 | -90.2355
1.5023 | -90.4867
1.5061 | -90.6917
1.5099 | -90.7738
1.5136 | -90.7868
1.5174 | -90.7644
1.5212 | -90.7591
1.525 | -90.8194
1.5287 | -90.9923
1.5325 | -91.2571
1.5363 | -91.4662
1.5401 | -91.6077
1.5438 | -91.7714
1.5476 | -91.6419
1.5514 | -91.0891
1.5552 | -90.5686
1.5589 | -90.3721
1.5627 | -90.5776
1.5665 | -91.122
1.5703 | -91.6816
1.574 | -91.8877
1.5778 | -91.5862
1.5816 | -90.9966
1.5854 | -90.6142
1.5891 | -90.7277
1.5929 | -91.1829
1.5967 | -91.538
1.6005 | -91.6448
1.6042 | -91.7554
1.608 | -91.9374
1.6118 | -92.0399
1.6156 | -91.9704
1.6193 | -91.8245
1.6231 | -91.8258
1.6269 | -91.955
1.6306 | -91.9042
1.6344 | -91.8033
1.6382 | -91.6643
1.642 | -91.5956
1.6457 | -91.6678
1.6495 | -91.8476
1.6533 | -91.9629
1.6571 | -91.9079
1.6608 | -91.7908
1.6646 | -91.7479
1.6684 | -91.7835
1.6722 | -91.8217
1.6759 | -91.8069
1.6797 | -91.8132
1.6835 | -91.9292
1.6873 | -92.1875
1.691 | -92.5007
1.6948 | -92.7127
1.6986 | -92.7836
1.7024 | -92.7788
1.7061 | -92.7627
1.7099 | -92.7266
1.7137 | -92.6061
1.7175 | -92.3998
1.7212 | -92.1422
1.725 | -91.9376
1.7288 | -91.9215
1.7326 | -92.1003
1.7363 | -92.3334
1.7401 | -92.4346
1.7439 | -92.4103
1.7477 | -92.3718
1.7514 | -92.359
1.7552 | -92.3628
1.759 | -92.3648
1.7628 | -92.3812
1.7665 | -92.4667
1.7703 | -92.6009
1.7741 | -92.716
1.7779 | -92.7466
1.7816 | -92.7562
1.7854 | -92.8413
1.7892 | -92.9958
1.793 | -93.0942
1.7967 | -93.0336
1.8005 | -92.8587
1.8043 | -92.6523
1.8081 | -92.461
1.8118 | -92.3067
1.8156 | -92.1796
1.8194 | -91.9905
1.8232 | -91.7115
1.8269 | -91.5031
1.8307 | -91.5101
1.8345 | -91.7303
1.8383 | -92.0581
1.842 | -92.4025
1.8458 | -92.7702
1.8496 | -93.1821
1.8534 | -93.5619
1.8571 | -93.7409
1.8609 | -93.6441
1.8647 | -93.411
1.8685 | -93.1661
1.8722 | -92.9527
1.876 | -92.767
1.8798 | -92.647
1.8836 | -92.6025
1.8873 | -92.6176
1.8911 | -92.6485
1.8949 | -92.665
1.8987 | -92.6456
1.9024 | -92.6173
1.9062 | -92.6338
1.91 | -92.7097
1.9137 | -92.7746
1.9175 | -92.768
1.9213 | -92.6876
1.9251 | -92.5752
1.9288 | -92.4456
1.9326 | -92.3432
1.9364 | -92.3347
1.9402 | -92.4437
1.9439 | -92.66
1.9477 | -92.9619
1.9515 | -93.3309
1.9553 | -93.6891
1.959 | -93.9264
1.9628 | -94.0349
1.9666 | -94.0866
1.9704 | -94.1574
1.9741 | -94.248
1.9779 | -94.3072
1.9817 | -94.3083
1.9855 | -94.2181
1.9892 | -94.0259
1.993 | -93.8093
1.9968 | -93.6648
2.0006 | -93.5826
2.0043 | -93.4717
2.0081 | -93.3245
2.0119 | -93.2267
2.0157 | -93.2419
2.0194 | -93.3154
2.0232 | -93.3428
2.027 | -93.2913
2.0308 | -93.2517
2.0345 | -93.2869
2.0383 | -93.3574
2.0421 | -93.374
2.0459 | -93.3574
2.0496 | -93.4231
2.0534 | -93.6324
2.0572 | -93.8744
2.061 | -93.9683
2.0647 | -93.8659
2.0685 | -93.6799
2.0723 | -93.5011
2.0761 | -93.4016
2.0798 | -93.4214
2.0836 | -93.5632
2.0874 | -93.7379
2.0912 | -93.8118
2.0949 | -93.7374
2.0987 | -93.6095
2.1025 | -93.5341
2.1063 | -93.5798
2.11 | -93.7309
2.1138 | -93.8905
2.1176 | -93.939
2.1214 | -93.8434
2.1251 | -93.6852
2.1289 | -93.5279
2.1327 | -93.3734
2.1365 | -93.2744
2.1402 | -93.339
2.144 | -93.6086
2.1478 | -93.954
2.1516 | -94.1986
2.1553 | -94.288
2.1591 | -94.2765
2.1629 | -94.198
2.1667 | -94.1087
2.1704 | -94.0701
2.1742 | -94.0654
2.178 | -94.0955
2.1818 | -94.1492
2.1855 | -94.2507
2.1893 | -94.3853
2.1931 | -94.5061
2.1968 | -94.5543
2.2006 | -94.5216
2.2044 | -94.4261
2.2082 | -94.3299
2.2119 | -94.2602
2.2157 | -94.1992
2.2195 | -94.1225
2.2233 | -94.0262
2.227 | -93.9118
2.2308 | -93.7864
2.2346 | -93.6962
2.2384 | -93.6783
2.2421 | -93.7655
2.2459 | -93.9692
2.2497 | -94.2327
2.2535 | -94.4829
2.2572 | -94.6488
2.261 | -94.7074
2.2648 | -94.699
2.2686 | -94.6749
2.2723 | -94.6479
2.2761 | -94.5856
2.2799 | -94.4781
2.2837 | -94.3331
2.2874 | -94.1997
2.2912 | -94.1466
2.295 | -94.1988
2.2988 | -94.3222
2.3025 | -94.4576
2.3063 | -94.5413
2.3101 | -94.5569
2.3139 | -94.5178
2.3176 | -94.4657
2.3214 | -94.408
2.3252 | -94.3604
2.329 | -94.3217
2.3327 | -94.3158
2.3365 | -94.3581
2.3403 | -94.4425
2.3441 | -94.5298
2.3478 | -94.5848
2.3516 | -94.5959
2.3554 | -94.5834
2.3592 | -94.5482
2.3629 | -94.5071
2.3667 | -94.4701
2.3705 | -94.4382
2.3743 | -94.4183
2.378 | -94.4082
2.3818 | -94.4197
2.3856 | -94.452
2.3894 | -94.4719
2.3931 | -94.4505
2.3969 | -94.4238
2.4007 | -94.4023
2.4045 | -94.3502
2.4082 | -94.2409
2.412 | -94.1003
2.4158 | -93.9684
2.4196 | -93.8941
2.4233 | -93.893
2.4271 | -93.9838
2.4309 | -94.1925
2.4347 | -94.5137
2.4384 | -94.8885
2.4422 | -95.272
2.446 | -95.558
2.4497 | -95.663
2.4535 | -95.5878
2.4573 | -95.3677
2.4611 | -95.068
2.4648 | -94.782
2.4686 | -94.6005
2.4724 | -94.5814
2.4762 | -94.7142
2.4799 | -94.9205
2.4837 | -95.0924
2.4875 | -95.1909
2.4913 | -95.2291
2.495 | -95.2777
2.4988 | -95.3773
2.5026 | -95.5051
2.5064 | -95.5953
2.5101 | -95.6087
2.5139 | -95.5564
2.5177 | -95.4795
2.5215 | -95.4014
2.5252 | -95.3328
2.529 | -95.2743
2.5328 | -95.2295
2.5366 | -95.2007
2.5403 | -95.1678
2.5441 | -95.1092
2.5479 | -95.023
2.5517 | -94.9491
2.5554 | -94.9452
2.5592 | -95.0785
2.563 | -95.3611
2.5668 | -95.7157
2.5705 | -96.0333
2.5743 | -96.2091
2.5781 | -96.199
2.5819 | -96.0362
2.5856 | -95.7999
2.5894 | -95.5766
2.5932 | -95.4195
2.597 | -95.3319
2.6007 | -95.2719
2.6045 | -95.2225
2.6083 | -95.1854
2.6121 | -95.1424
2.6158 | -95.107
2.6196 | -95.1143
2.6234 | -95.2101
2.6272 | -95.3815
2.6309 | -95.5939
2.6347 | -95.7736
2.6385 | -95.864
2.6423 | -95.8528
2.646 | -95.7738
2.6498 | -95.6745
2.6536 | -95.604
2.6574 | -95.6047
2.6611 | -95.7103
2.6649 | -95.9303
2.6687 | -96.2137
2.6725 | -96.4322
2.6762 | -96.4364
2.68 | -96.2125
2.6838 | -95.8623
2.6876 | -95.5058
2.6913 | -95.2455
2.6951 | -95.127
2.6989 | -95.1748
2.7027 | -95.3809
2.7064 | -95.689
2.7102 | -96.006
2.714 | -96.2414
2.7178 | -96.3373
2.7215 | -96.1695
2.7253 | -96.0723
2.7291 | -95.9535
2.7328 | -95.866
2.7366 | -95.8406
2.7404 | -95.8796
2.7442 | -95.9432
2.7479 | -96.0075
2.7517 | -96.0549
2.7555 | -96.0924
2.7593 | -96.1172
2.763 | -96.1314
2.7668 | -96.1219
2.7706 | -96.0808
2.7744 | -96.0358
2.7781 | -96.0173
2.7819 | -96.0449
2.7857 | -96.112
2.7895 | -96.2125
2.7932 | -96.3289
2.797 | -96.4232
2.8008 | -96.471
2.8046 | -96.4682
2.8083 | -96.4239
2.8121 | -96.3892
2.8159 | -96.3793
2.8197 | -96.3972
2.8234 | -96.4097
2.8272 | -96.4012
2.831 | -96.3542
2.8348 | -96.2848
2.8385 | -96.2236
2.8423 | -96.2069
2.8461 | -96.2442
2.8499 | -96.298
2.8536 | -96.3362
2.8574 | -96.3298
2.8612 | -96.2576
2.865 | -96.1555
2.8687 | -96.0334
2.8725 | -95.9278
2.8763 | -95.8688
2.8801 | -95.8815
2.8838 | -95.968
2.8876 | -96.0971
2.8914 | -96.2195
2.8952 | -96.2857
2.8989 | -96.281
2.9027 | -96.198
2.9065 | -96.0812
2.9103 | -95.9664
2.914 | -95.883
2.9178 | -95.8517
2.9216 | -95.8644
2.9254 | -95.8904
2.9291 | -95.919
2.9329 | -95.9407
2.9367 | -95.9806
2.9405 | -96.054
2.9442 | -96.1591
2.948 | -96.2685
2.9518 | -96.351
2.9556 | -96.3794
2.9593 | -96.3526
2.9631 | -96.311
2.9669 | -96.2959
2.9707 | -96.3411
2.9744 | -96.4568
2.9782 | -96.6052
2.982 | -96.7499
2.9858 | -96.8411
2.9895 | -96.8527
2.9933 | -96.7927
2.9971 | -96.671
3.0008 | -96.5296
3.0046 | -96.3922
3.0084 | -96.2981
3.0122 | -96.2537
3.0159 | -96.2663
3.0197 | -96.3387
3.0235 | -96.4838
3.0273 | -96.6855
3.031 | -96.9046
3.0348 | -97.103
3.0386 | -97.217
3.0424 | -97.2202
3.0461 | -97.1472
3.0499 | -97.0484
3.0537 | -96.9701
3.0575 | -96.9338
3.0612 | -96.9563
3.065 | -97.0288
3.0688 | -97.1183
3.0726 | -97.1858
3.0763 | -97.2106
3.0801 | -97.1834
3.0839 | -97.1087
3.0877 | -97.0132
3.0914 | -96.9218
3.0952 | -96.8546
3.099 | -96.8055
3.1028 | -96.7873
3.1065 | -96.7757
3.1103 | -96.7549
3.1141 | -96.7424
3.1179 | -96.7494
3.1216 | -96.787
3.1254 | -96.8505
3.1292 | -96.9061
3.133 | -96.9371
3.1367 | -96.9429
3.1405 | -96.9266
3.1443 | -96.9218
3.1481 | -96.9604
3.1518 | -97.0232
3.1556 | -97.0914
3.1594 | -97.1601
3.1632 | -97.224
3.1669 | -97.29
3.1707 | -97.3587
3.1745 | -97.4142
3.1783 | -97.4383
3.182 | -97.4218
3.1858 | -97.36
3.1896 | -97.2625
3.1934 | -97.1298
3.1971 | -96.9904
3.2009 | -96.8626
3.2047 | -96.76
3.2085 | -96.6923
3.2122 | -96.672
3.216 | -96.7109
3.2198 | -96.7968
3.2236 | -96.9106
3.2273 | -97.034
3.2311 | -97.1124
3.2349 | -97.134
3.2387 | -97.0949
3.2424 | -97.0141
3.2462 | -96.9026
3.25 | -96.8118
3.2538 | -96.7738
3.2575 | -96.7858
3.2613 | -96.8322
3.2651 | -96.8684
3.2689 | -96.9096
3.2726 | -96.9541
3.2764 | -97.0027
3.2802 | -97.0529
3.2839 | -97.1181
3.2877 | -97.1789
3.2915 | -97.2048
3.2953 | -97.188
3.299 | -97.1331
3.3028 | -97.0414
3.3066 | -96.9237
3.3104 | -96.8187
3.3141 | -96.7511
3.3179 | -96.7371
3.3217 | -96.7629
3.3255 | -96.8162
3.3292 | -96.8958
3.333 | -96.9914
3.3368 | -97.0957
3.3406 | -97.1906
3.3443 | -97.2673
3.3481 | -97.3131
3.3519 | -97.3226
3.3557 | -97.3175
3.3594 | -97.3118
3.3632 | -97.3268
3.367 | -97.341
3.3708 | -97.3373
3.3745 | -97.3064
3.3783 | -97.2319
3.3821 | -97.121
3.3859 | -97.0002
3.3896 | -96.8995
3.3934 | -96.8436
3.3972 | -96.8596
3.401 | -96.9407
3.4047 | -97.0769
3.4085 | -97.2355
3.4123 | -97.3853
3.4161 | -97.5153
3.4198 | -97.6192
3.4236 | -97.7134
3.4274 | -97.8088
3.4312 | -97.8878
3.4349 | -97.9592
3.4387 | -98.0129
3.4425 | -98.0176
3.4463 | -97.983
3.45 | -97.9187
3.4538 | -97.8323
3.4576 | -97.7361
3.4614 | -97.6353
3.4651 | -97.5471
3.4689 | -97.4857
3.4727 | -97.4518
3.4765 | -97.4398
3.4802 | -97.4436
3.484 | -97.4549
3.4878 | -97.4798
3.4916 | -97.505
3.4953 | -97.525
3.4991 | -97.5529
3.5029 | -97.5818
3.5067 | -97.6135
3.5104 | -97.6352
3.5142 | -97.6451
3.518 | -97.6364
3.5218 | -97.6053
3.5255 | -97.5588
3.5293 | -97.5184
3.5331 | -97.5011
3.5368 | -97.5223
3.5406 | -97.5764
3.5444 | -97.6472
3.5482 | -97.7191
3.5519 | -97.7624
3.5557 | -97.7662
3.5595 | -97.7261
3.5633 | -97.659
3.567 | -97.5965
3.5708 | -97.5601
3.5746 | -97.5492
3.5784 | -97.5548
3.5821 | -97.573
3.5859 | -97.5945
3.5897 | -97.6081
3.5935 | -97.6158
3.5972 | -97.6404
3.601 | -97.6816
3.6048 | -97.7402
3.6086 | -97.8277
3.6123 | -97.9186
3.6161 | -98.0067
3.6199 | -98.07
3.6237 | -98.1079
3.6274 | -98.1341
3.6312 | -98.1664
3.635 | -98.199
3.6388 | -98.2224
3.6425 | -98.2297
3.6463 | -98.2123
3.6501 | -98.161
3.6539 | -98.0641
3.6576 | -97.9592
3.6614 | -97.8587
3.6652 | -97.7773
3.669 | -97.7455
3.6727 | -97.7443
3.6765 | -97.7751
3.6803 | -97.8004
3.6841 | -97.8175
3.6878 | -97.8329
3.6916 | -97.8535
3.6954 | -97.8875
3.6992 | -97.9322
3.7029 | -97.9909
3.7067 | -98.0533
3.7105 | -98.1259
3.7143 | -98.1868
3.718 | -98.2545
3.7218 | -98.3251
3.7256 | -98.3894
3.7294 | -98.4513
3.7331 | -98.4827
3.7369 | -98.4771
3.7407 | -98.4326
3.7445 | -98.3531
3.7482 | -98.2542
3.752 | -98.1473
3.7558 | -98.057
3.7596 | -98.0021
3.7633 | -98.009
3.7671 | -98.072
3.7709 | -98.1913
3.7747 | -98.343
3.7784 | -98.5037
3.7822 | -98.6507
3.786 | -98.776
3.7898 | -98.8701
3.7935 | -98.9204
3.7973 | -98.9084
3.8011 | -98.8415
3.8049 | -98.7223
3.8086 | -98.5389
3.8124 | -98.3828
3.8162 | -98.229
3.8199 | -98.0869
3.8237 | -97.9598
3.8275 | -97.863
3.8313 | -97.8134
3.835 | -97.8129
3.8388 | -97.8651
3.8426 | -97.962
3.8464 | -98.0845
3.8501 | -98.2208
3.8539 | -98.3489
3.8577 | -98.444
3.8615 | -98.4994
3.8652 | -98.5157
3.869 | -98.5059
3.8728 | -98.4829
3.8766 | -98.4588
3.8803 | -98.4494
3.8841 | -98.4624
3.8879 | -98.5054
3.8917 | -98.5794
3.8954 | -98.6587
3.8992 | -98.7484
3.903 | -98.8323
3.9068 | -98.8993
3.9105 | -98.9554
3.9143 | -99.0003
3.9181 | -99.0366
3.9219 | -99.0597
3.9256 | -99.081
3.9294 | -99.0918
3.9332 | -99.0843
3.937 | -99.0737
3.9407 | -99.0787
3.9445 | -99.0892
3.9483 | -99.1153
3.9521 | -99.1441
3.9558 | -99.1654
3.9596 | -99.1775
3.9634 | -99.1653
3.9672 | -99.1299
3.9709 | -99.0801
3.9747 | -99.0332
3.9785 | -98.9981
3.9823 | -98.9868
3.986 | -98.9937
3.9898 | -98.999
3.9936 | -99.0001
3.9974 | -98.9946
4.0011 | -98.9729
4.0049 | -98.935
4.0087 | -98.8912
4.0125 | -98.8561
4.0162 | -98.8376
4.02 | -98.8316
4.0238 | -98.8451
4.0276 | -98.8693
4.0313 | -98.8945
4.0351 | -98.9064
4.0389 | -98.8901
4.0427 | -98.8529
4.0464 | -98.7997
4.0502 | -98.752
4.054 | -98.7158
4.0578 | -98.7019
4.0615 | -98.7129
4.0653 | -98.7492
4.0691 | -98.797
4.0729 | -98.8478
4.0766 | -98.8931
4.0804 | -98.9443
4.0842 | -99.0064
4.0879 | -99.0608
4.0917 | -99.1184
4.0955 | -99.1773
4.0993 | -99.2179
4.103 | -99.2443
4.1068 | -99.2585
4.1106 | -99.2509
4.1144 | -99.2193
4.1181 | -99.1917
4.1219 | -99.1701
4.1257 | -99.143
4.1295 | -99.107
4.1332 | -99.0726
4.137 | -99.0357
4.1408 | -98.9922
4.1446 | -98.9542
4.1483 | -98.9289
4.1521 | -98.9292
4.1559 | -98.9438
4.1597 | -98.9647
4.1634 | -99.0021
4.1672 | -99.0378
4.171 | -99.0712
4.1748 | -99.0971
4.1785 | -99.1183
4.1823 | -99.1076
4.1861 | -99.0765
4.1899 | -99.032
4.1936 | -98.9843
4.1974 | -98.9295
4.2012 | -98.8843
4.205 | -98.859
4.2087 | -98.8605
4.2125 | -98.8918
4.2163 | -98.951
4.2201 | -99.0293
4.2238 | -99.1218
4.2276 | -99.2084
4.2314 | -99.2892
4.2352 | -99.3615
4.2389 | -99.4308
4.2427 | -99.4875
4.2465 | -99.5375
4.2503 | -99.5688
4.254 | -99.5776
4.2578 | -99.5606
4.2616 | -99.5238
4.2654 | -99.4749
4.2691 | -99.4228
4.2729 | -99.3637
4.2767 | -99.2982
4.2805 | -99.2404
4.2842 | -99.2019
4.288 | -99.1797
4.2918 | -99.1883
4.2956 | -99.2176
4.2993 | -99.2531
4.3031 | -99.2973
4.3069 | -99.3369
4.3107 | -99.3756
4.3144 | -99.423
4.3182 | -99.481
4.322 | -99.5442
4.3258 | -99.6201
4.3295 | -99.6928
4.3333 | -99.7629
4.3371 | -99.8328
4.3409 | -99.8946
4.3446 | -99.9333
4.3484 | -99.9588
4.3522 | -99.9637
4.3559 | -99.9119
4.3597 | -99.8675
4.3635 | -99.8103
4.3673 | -99.751
4.371 | -99.6819
4.3748 | -99.6171
4.3786 | -99.5622
4.3824 | -99.5151
4.3861 | -99.4721
4.3899 | -99.4375
4.3937 | -99.4027
4.3975 | -99.3677
4.4012 | -99.3422
4.405 | -99.3234
4.4088 | -99.3213
4.4126 | -99.3292
4.4163 | -99.3458
4.4201 | -99.3828
4.4239 | -99.4364
4.4277 | -99.5061
4.4314 | -99.57
4.4352 | -99.6146
4.439 | -99.6275
4.4428 | -99.6141
4.4465 | -99.5715
4.4503 | -99.5048
4.4541 | -99.4313
4.4579 | -99.3629
4.4616 | -99.3098
4.4654 | -99.2883
4.4692 | -99.2976
4.473 | -99.3258
4.4767 | -99.3755
4.4805 | -99.4411
4.4843 | -99.5146
4.4881 | -99.5836
4.4918 | -99.6479
4.4956 | -99.7005
4.4994 | -99.7477
4.5032 | -99.8021
4.5069 | -99.8574
4.5107 | -99.9164
4.5145 | -99.9748
4.5183 | -100.0264
4.522 | -100.0721
4.5258 | -100.1008
4.5296 | -100.1235
4.5334 | -100.1425
4.5371 | -100.1541
4.5409 | -100.1653
4.5447 | -100.1747
4.5485 | -100.191
4.5522 | -100.2056
4.556 | -100.2192
4.5598 | -100.2297
4.5636 | -100.222
4.5673 | -100.2034
4.5711 | -100.1768
4.5749 | -100.1337
4.5787 | -100.072
4.5824 | -99.9965
4.5862 | -99.918
4.59 | -99.8483
4.5938 | -99.7818
4.5975 | -99.737
4.6013 | -99.7085
4.6051 | -99.6898
4.6089 | -99.6807
4.6126 | -99.676
4.6164 | -99.6696
4.6202 | -99.6535
4.6239 | -99.6375
4.6277 | -99.6316
4.6315 | -99.6401
4.6353 | -99.6755
4.639 | -99.724
4.6428 | -99.7813
4.6466 | -99.8475
4.6504 | -99.9065
4.6541 | -99.9504
4.6579 | -99.9736
4.6617 | -99.9772
4.6655 | -99.9671
4.6692 | -99.9445
4.673 | -99.9168
4.6768 | -99.8931
4.6806 | -99.8785
4.6843 | -99.8668
4.6881 | -99.8602
4.6919 | -99.855
4.6957 | -99.8506
4.6994 | -99.8501
4.7032 | -99.8583
4.707 | -99.8766
4.7108 | -99.9011
4.7145 | -99.931
4.7183 | -99.9634
4.7221 | -99.9951
4.7259 | -100.0224
4.7296 | -100.0423
4.7334 | -100.0415
4.7372 | -100.0306
4.741 | -100.001
4.7447 | -99.9574
4.7485 | -99.9042
4.7523 | -99.85
4.7561 | -99.8095
4.7598 | -99.7824
4.7636 | -99.7862
4.7674 | -99.8259
4.7712 | -99.8854
4.7749 | -99.9562
4.7787 | -100.0194
4.7825 | -100.071
4.7863 | -100.101
4.79 | -100.1119
4.7938 | -100.1039
4.7976 | -100.0775
4.8014 | -100.0376
4.8051 | -99.9968
4.8089 | -99.9413
4.8127 | -99.8941
4.8165 | -99.8556
4.8202 | -99.8395
4.824 | -99.8515
4.8278 | -99.8752
4.8316 | -99.9026
4.8353 | -99.9147
4.8391 | -99.9195
4.8429 | -99.9265
4.8467 | -99.9364
4.8504 | -99.9599
4.8542 | -100.0076
4.858 | -100.0698
4.8618 | -100.121
4.8655 | -100.1671
4.8693 | -100.1921
4.8731 | -100.1857
4.8769 | -100.1447
4.8806 | -100.0942
4.8844 | -100.0379
4.8882 | -99.9951
4.892 | -99.9652
4.8957 | -99.9527
4.8995 | -99.9987
4.9033 | -100.0258
4.907 | -100.0649
4.9108 | -100.111
4.9146 | -100.148
4.9184 | -100.1878
4.9221 | -100.2262
4.9259 | -100.2655
4.9297 | -100.3011
4.9335 | -100.3412
4.9372 | -100.4019
4.941 | -100.4643
4.9448 | -100.5323
4.9486 | -100.6029
4.9523 | -100.6658
4.9561 | -100.7142
4.9599 | -100.7372
4.9637 | -100.7519
4.9674 | -100.7445
4.9712 | -100.7178
4.975 | -100.6739
4.9788 | -100.6095
4.9825 | -100.5342
4.9863 | -100.4584
4.9901 | -100.3906
4.9939 | -100.3319
4.9976 | -100.2851
5.0014 | -100.2605
5.0052 | -100.2552
5.009 | -100.2693
5.0127 | -100.2925
5.0165 | -100.321
5.0203 | -100.3486
5.0241 | -100.3704
5.0278 | -100.3821
5.0316 | -100.3866
5.0354 | -100.3737
5.0392 | -100.3576
5.0429 | -100.3298
5.0467 | -100.3176
5.0505 | -100.319
5.0543 | -100.337
5.058 | -100.3769
5.0618 | -100.4382
5.0656 | -100.5184
5.0694 | -100.6092
5.0731 | -100.7102
5.0769 | -100.8186
5.0807 | -100.9301
5.0845 | -101.0317
5.0882 | -101.1261
5.092 | -101.2203
5.0958 | -101.2988
5.0996 | -101.3582
5.1033 | -101.3949
5.1071 | -101.4172
5.1109 | -101.4204
5.1147 | -101.4135
5.1184 | -101.3903
5.1222 | -101.3539
5.126 | -101.3012
5.1298 | -101.231
5.1335 | -101.1389
5.1373 | -101.0358
5.1411 | -100.9302
5.1449 | -100.8263
5.1486 | -100.7332
5.1524 | -100.6533
5.1562 | -100.589
5.16 | -100.5414
5.1637 | -100.5163
5.1675 | -100.5079
5.1713 | -100.5053
5.175 | -100.5112
5.1788 | -100.5275
5.1826 | -100.5451
5.1864 | -100.5643
5.1901 | -100.5828
5.1939 | -100.6005
5.1977 | -100.6193
5.2015 | -100.6351
5.2052 | -100.6483
5.209 | -100.6705
5.2128 | -100.7024
5.2166 | -100.7361
5.2203 | -100.7753
5.2241 | -100.8233
5.2279 | -100.8703
5.2317 | -100.9219
5.2354 | -100.954
5.2392 | -100.9708
5.243 | -100.9815
5.2468 | -100.9709
5.2505 | -100.9557
5.2543 | -100.9401
5.2581 | -100.9208
5.2619 | -100.9057
5.2656 | -100.8931
5.2694 | -100.8842
5.2732 | -100.8847
5.277 | -100.8903
5.2807 | -100.8967
5.2845 | -100.9062
5.2883 | -100.9204
5.2921 | -100.949
5.2958 | -100.9899
5.2996 | -101.0411
5.3034 | -101.1018
5.3072 | -101.1746
5.3109 | -101.2361
5.3147 | -101.2995
5.3185 | -101.3554
5.3223 | -101.3879
5.326 | -101.4029
5.3298 | -101.3922
5.3336 | -101.3521
5.3374 | -101.2995
5.3411 | -101.2345
5.3449 | -101.1606
5.3487 | -101.0942
5.3525 | -101.0402
5.3562 | -101.0135
5.36 | -101.0176
5.3638 | -101.0489
5.3676 | -101.1005
5.3713 | -101.1715
5.3751 | -101.2475
5.3789 | -101.3234
5.3827 | -101.3921
5.3864 | -101.4367
5.3902 | -101.4611
5.394 | -101.448
5.3978 | -101.3878
5.4015 | -101.2972
5.4053 | -101.1853
5.4091 | -101.0616
5.4129 | -100.9414
5.4166 | -100.8309
5.4204 | -100.7445
5.4242 | -100.6876
5.428 | -100.655
5.4317 | -100.6492
5.4355 | -100.6709
5.4393 | -100.7169
5.4431 | -100.7959
5.4468 | -100.8582
5.4506 | -100.9324
5.4544 | -101.0046
5.4581 | -101.0783
5.4619 | -101.1432
5.4657 | -101.192
5.4695 | -101.2369
5.4732 | -101.2693
5.477 | -101.2855
5.4808 | -101.2981
5.4846 | -101.2968
5.4883 | -101.2796
5.4921 | -101.2393
5.4959 | -101.1775
5.4997 | -101.1002
5.5034 | -101.0155
5.5072 | -100.9303
5.511 | -100.8498
5.5148 | -100.7756
5.5185 | -100.721
5.5223 | -100.6854
5.5261 | -100.672
5.5299 | -100.6781
5.5336 | -100.7111
5.5374 | -100.7626
5.5412 | -100.8277
5.545 | -100.907
5.5487 | -100.9968
5.5525 | -101.0865
5.5563 | -101.1728
5.5601 | -101.2562
5.5638 | -101.3383
5.5676 | -101.415
5.5714 | -101.479
5.5752 | -101.5293
5.5789 | -101.5784
5.5827 | -101.615
5.5865 | -101.628
5.5903 | -101.6306
5.594 | -101.6104
5.5978 | -101.575
5.6016 | -101.5259
5.6054 | -101.459
5.6091 | -101.3819
5.6129 | -101.3008
5.6167 | -101.2146
5.6205 | -101.1282
5.6242 | -101.0573
5.628 | -101.0095
5.6318 | -100.9825
5.6356 | -100.9744
5.6393 | -100.9887
5.6431 | -101.0177
5.6469 | -101.0568
5.6507 | -101.0966
5.6544 | -101.1323
5.6582 | -101.1611
5.662 | -101.1806
5.6658 | -101.187
5.6695 | -101.1934
5.6733 | -101.2002
5.6771 | -101.1945
5.6809 | -101.1969
5.6846 | -101.1892
5.6884 | -101.1788
5.6922 | -101.1749
5.696 | -101.1774
5.6997 | -101.1763
5.7035 | -101.1745
5.7073 | -101.1749
5.7111 | -101.1775
5.7148 | -101.1764
5.7186 | -101.16
5.7224 | -101.135
5.7261 | -101.1091
5.7299 | -101.0791
5.7337 | -101.0575
5.7375 | -101.0507
5.7412 | -101.0574
5.745 | -101.0858
5.7488 | -101.1356
5.7526 | -101.2048
5.7563 | -101.2831
5.7601 | -101.3819
5.7639 | -101.4935
5.7677 | -101.6013
5.7714 | -101.7133
5.7752 | -101.8062
5.779 | -101.8719
5.7828 | -101.9162
5.7865 | -101.9464
5.7903 | -101.9615
5.7941 | -101.9545
5.7979 | -101.9402
5.8016 | -101.9275
5.8054 | -101.9255
5.8092 | -101.931
5.813 | -101.9348
5.8167 | -101.9358
5.8205 | -101.9322
5.8243 | -101.9228
5.8281 | -101.9055
5.8318 | -101.8811
5.8356 | -101.8447
5.8394 | -101.7992
5.8432 | -101.7609
5.8469 | -101.7243
5.8507 | -101.6902
5.8545 | -101.6665
5.8583 | -101.6512
5.862 | -101.641
5.8658 | -101.6333
5.8696 | -101.6265
5.8734 | -101.6134
5.8771 | -101.5963
5.8809 | -101.5792
5.8847 | -101.5619
5.8885 | -101.5482
5.8922 | -101.5329
5.896 | -101.5211
5.8998 | -101.5189
5.9036 | -101.5303
5.9073 | -101.5453
5.9111 | -101.5663
5.9149 | -101.5944
5.9187 | -101.6336
5.9224 | -101.6754
5.9262 | -101.716
5.93 | -101.7516
5.9338 | -101.7834
5.9375 | -101.8107
5.9413 | -101.8231
5.9451 | -101.8236
5.9489 | -101.82
5.9526 | -101.8058
5.9564 | -101.7911
5.9602 | -101.7763
5.964 | -101.7613
5.9677 | -101.7475
5.9715 | -101.7353
5.9753 | -101.7321
5.9791 | -101.7418
5.9828 | -101.7444
5.9866 | -101.7292
5.9904 | -101.7202
5.9941 | -101.7052
5.9979 | -101.6882
6.0017 | -101.6655
6.0055 | -101.637
6.0092 | -101.6022
6.013 | -101.5741
6.0168 | -101.5513
6.0206 | -101.5347
6.0243 | -101.5302
6.0281 | -101.5434
6.0319 | -101.5751
6.0357 | -101.6195
6.0394 | -101.6752
6.0432 | -101.7476
6.047 | -101.829
6.0508 | -101.9106
6.0545 | -101.9954
6.0583 | -102.0812
6.0621 | -102.1574
6.0659 | -102.2233
6.0696 | -102.2729
6.0734 | -102.3099
6.0772 | -102.3268
6.081 | -102.3127
6.0847 | -102.2787
6.0885 | -102.2286
6.0923 | -102.1583
6.0961 | -102.0785
6.0998 | -101.991
6.1036 | -101.902
6.1074 | -101.8182
6.1112 | -101.7456
6.1149 | -101.6874
6.1187 | -101.6406
6.1225 | -101.6076
6.1263 | -101.5866
6.13 | -101.5758
6.1338 | -101.5688
6.1376 | -101.5652
6.1414 | -101.5599
6.1451 | -101.5667
6.1489 | -101.5654
6.1527 | -101.5585
6.1565 | -101.5516
6.1602 | -101.5441
6.164 | -101.5368
6.1678 | -101.5315
6.1716 | -101.5292
6.1753 | -101.5272
6.1791 | -101.5338
6.1829 | -101.5523
6.1867 | -101.5775
6.1904 | -101.6031
6.1942 | -101.6262
6.198 | -101.6465
6.2018 | -101.6631
6.2055 | -101.6761
6.2093 | -101.6856
6.2131 | -101.6907
6.2169 | -101.6906
6.2206 | -101.6841
6.2244 | -101.6722
6.2282 | -101.6565
6.232 | -101.6363
6.2357 | -101.6195
6.2395 | -101.6031
6.2433 | -101.5975
6.2471 | -101.5993
6.2508 | -101.6078
6.2546 | -101.6287
6.2584 | -101.6621
6.2622 | -101.703
6.2659 | -101.7522
6.2697 | -101.8045
6.2735 | -101.8523
6.2772 | -101.9016
6.281 | -101.9571
6.2848 | -102.0116
6.2886 | -102.0681
6.2923 | -102.1163
6.2961 | -102.151
6.2999 | -102.1734
6.3037 | -102.1884
6.3074 | -102.1907
6.3112 | -102.1952
6.315 | -102.2007
6.3188 | -102.2059
6.3225 | -102.2079
6.3263 | -102.2093
6.3301 | -102.2156
6.3339 | -102.2282
6.3376 | -102.2409
6.3414 | -102.2579
6.3452 | -102.2796
6.349 | -102.2973
6.3527 | -102.3189
6.3565 | -102.3418
6.3603 | -102.3642
6.3641 | -102.3918
6.3678 | -102.4178
6.3716 | -102.4437
6.3754 | -102.4734
6.3792 | -102.5113
6.3829 | -102.5439
6.3867 | -102.5743
6.3905 | -102.6013
6.3943 | -102.6187
6.398 | -102.6266
6.4018 | -102.6216
6.4056 | -102.5936
6.4094 | -102.5557
6.4131 | -102.5118
6.4169 | -102.4671
6.4207 | -102.42
6.4245 | -102.3727
6.4282 | -102.3371
6.432 | -102.3232
6.4358 | -102.3274
6.4396 | -102.3411
6.4433 | -102.3693
6.4471 | -102.407
6.4509 | -102.4547
6.4547 | -102.5072
6.4584 | -102.5584
6.4622 | -102.6116
6.466 | -102.6589
6.4698 | -102.6916
6.4735 | -102.7159
6.4773 | -102.7349
6.4811 | -102.7388
6.4849 | -102.7279
6.4886 | -102.7003
6.4924 | -102.6539
6.4962 | -102.6031
6.5 | -102.5424
6.5037 | -102.4739
6.5075 | -102.4056
6.5113 | -102.3337
6.5151 | -102.2627
6.5188 | -102.2006
6.5226 | -102.1411
6.5264 | -102.0819
6.5301 | -102.0333
6.5339 | -102.0146
6.5377 | -101.9876
6.5415 | -101.9726
6.5452 | -101.9686
6.549 | -101.9734
6.5528 | -101.9822
6.5566 | -102.0051
6.5603 | -102.0358
6.5641 | -102.0695
6.5679 | -102.115
6.5717 | -102.1609
6.5754 | -102.2105
6.5792 | -102.2693
6.583 | -102.3249
6.5868 | -102.3847
6.5905 | -102.4438
6.5943 | -102.5024
6.5981 | -102.5618
6.6019 | -102.6196
6.6056 | -102.6728
6.6094 | -102.7272
6.6132 | -102.7789
6.617 | -102.8202
6.6207 | -102.859
6.6245 | -102.8938
6.6283 | -102.9127
6.6321 | -102.9246
6.6358 | -102.9294
6.6396 | -102.9246
6.6434 | -102.9098
6.6472 | -102.8951
6.6509 | -102.8795
6.6547 | -102.8658
6.6585 | -102.8469
6.6623 | -102.8263
6.666 | -102.8113
6.6698 | -102.7906
6.6736 | -102.7699
6.6774 | -102.7472
6.6811 | -102.7189
6.6849 | -102.6943
6.6887 | -102.6663
6.6925 | -102.6402
6.6962 | -102.6126
6.7 | -102.5835
6.7038 | -102.5528
6.7076 | -102.5197
6.7113 | -102.4911
6.7151 | -102.4695
6.7189 | -102.4455
6.7227 | -102.4227
6.7264 | -102.4027
6.7302 | -102.3856
6.734 | -102.3718
6.7378 | -102.3599
6.7415 | -102.355
6.7453 | -102.3527
6.7491 | -102.3559
6.7529 | -102.368
6.7566 | -102.3817
6.7604 | -102.395
6.7642 | -102.4165
6.768 | -102.4402
6.7717 | -102.4712
6.7755 | -102.5077
6.7793 | -102.5391
6.7831 | -102.5657
6.7868 | -102.5894
6.7906 | -102.6069
6.7944 | -102.6181
6.7981 | -102.6225
6.8019 | -102.6204
6.8057 | -102.6162
6.8095 | -102.6043
6.8132 | -102.5855
6.817 | -102.5641
6.8208 | -102.5558
6.8246 | -102.551
6.8283 | -102.5505
6.8321 | -102.555
6.8359 | -102.5648
6.8397 | -102.58
6.8434 | -102.5946
6.8472 | -102.6114
6.851 | -102.6265
6.8548 | -102.6409
6.8585 | -102.6552
6.8623 | -102.6701
6.8661 | -102.6766
6.8699 | -102.6747
6.8736 | -102.661
6.8774 | -102.6399
6.8812 | -102.6193
6.885 | -102.5976
6.8887 | -102.5671
6.8925 | -102.535
6.8963 | -102.5105
6.9001 | -102.4955
6.9038 | -102.4888
6.9076 | -102.4943
6.9114 | -102.5131
6.9152 | -102.5427
6.9189 | -102.581
6.9227 | -102.6296
6.9265 | -102.6966
6.9303 | -102.7716
6.934 | -102.8526
6.9378 | -102.9295
6.9416 | -103.005
6.9454 | -103.082
6.9491 | -103.1464
6.9529 | -103.1886
6.9567 | -103.2122
6.9605 | -103.2161
6.9642 | -103.194
6.968 | -103.1551
6.9718 | -103.1077
6.9756 | -103.0482
6.9793 | -102.9801
6.9831 | -102.9168
6.9869 | -102.8533
6.9907 | -102.7987
6.9944 | -102.7489
6.9982 | -102.701
7.002 | -102.6583
7.0058 | -102.6281
7.0095 | -102.6034
7.0133 | -102.5918
7.0171 | -102.5891
7.0209 | -102.6028
7.0246 | -102.6264
7.0284 | -102.6529
7.0322 | -102.6947
7.036 | -102.7464
7.0397 | -102.8069
7.0435 | -102.8735
7.0473 | -102.9389
7.0511 | -103.0124
7.0548 | -103.0869
7.0586 | -103.1601
7.0624 | -103.2314
7.0662 | -103.3011
7.0699 | -103.3587
7.0737 | -103.4102
7.0775 | -103.4366
7.0812 | -103.4722
7.085 | -103.5004
7.0888 | -103.5112
7.0926 | -103.5132
7.0963 | -103.5058
7.1001 | -103.4878
7.1039 | -103.4589
7.1077 | -103.4349
7.1114 | -103.4041
7.1152 | -103.3746
7.119 | -103.3447
7.1228 | -103.3209
7.1265 | -103.2971
7.1303 | -103.2803
7.1341 | -103.2715
7.1379 | -103.2623
7.1416 | -103.2543
7.1454 | -103.2464
7.1492 | -103.2369
7.153 | -103.226
7.1567 | -103.2116
7.1605 | -103.1952
7.1643 | -103.1749
7.1681 | -103.1515
7.1718 | -103.127
7.1756 | -103.1013
7.1794 | -103.0677
7.1832 | -103.0321
7.1869 | -102.999
7.1907 | -102.9709
7.1945 | -102.9435
7.1983 | -102.9121
7.202 | -102.8845
7.2058 | -102.8639
7.2096 | -102.8502
7.2134 | -102.84
7.2171 | -102.8187
7.2209 | -102.8035
7.2247 | -102.7897
7.2285 | -102.773
7.2322 | -102.7495
7.236 | -102.7198
7.2398 | -102.6852
7.2436 | -102.6502
7.2473 | -102.617
7.2511 | -102.586
7.2549 | -102.5605
7.2587 | -102.5387
7.2624 | -102.5216
7.2662 | -102.5112
7.27 | -102.5092
7.2738 | -102.5138
7.2775 | -102.5251
7.2813 | -102.5368
7.2851 | -102.5578
7.2889 | -102.5776
7.2926 | -102.5972
7.2964 | -102.6193
7.3002 | -102.6396
7.304 | -102.6564
7.3077 | -102.6719
7.3115 | -102.6784
7.3153 | -102.6732
7.3191 | -102.6615
7.3228 | -102.6463
7.3266 | -102.6359
7.3304 | -102.6236
7.3342 | -102.6117
7.3379 | -102.608
7.3417 | -102.6133
7.3455 | -102.6285
7.3492 | -102.6509
7.353 | -102.6765
7.3568 | -102.7069
7.3606 | -102.7448
7.3643 | -102.7801
7.3681 | -102.8136
7.3719 | -102.8512
7.3757 | -102.8864
7.3794 | -102.917
7.3832 | -102.9445
7.387 | -102.9735
7.3908 | -103.0052
7.3945 | -103.0346
7.3983 | -103.074
7.4021 | -103.1069
7.4059 | -103.1408
7.4096 | -103.1776
7.4134 | -103.2184
7.4172 | -103.2668
7.421 | -103.3137
7.4247 | -103.3523
7.4285 | -103.3878
7.4323 | -103.4166
7.4361 | -103.4338
7.4398 | -103.4404
7.4436 | -103.4387
7.4474 | -103.4211
7.4512 | -103.3942
7.4549 | -103.3525
7.4587 | -103.3038
7.4625 | -103.256
7.4663 | -103.2045
7.47 | -103.1483
7.4738 | -103.1004
7.4776 | -103.055
7.4814 | -103.0106
7.4851 | -102.9726
7.4889 | -102.9417
7.4927 | -102.9216
7.4965 | -102.9097
7.5002 | -102.9087
7.504 | -102.9219
7.5078 | -102.9416
7.5116 | -102.9799
7.5153 | -103.0255
7.5191 | -103.0731
7.5229 | -103.119
7.5267 | -103.1624
7.5304 | -103.204
7.5342 | -103.2445
7.538 | -103.2709
7.5418 | -103.2916
7.5455 | -103.3061
7.5493 | -103.3143
7.5531 | -103.3203
7.5569 | -103.3307
7.5606 | -103.342
7.5644 | -103.3557
7.5682 | -103.3689
7.572 | -103.3786
7.5757 | -103.3913
7.5795 | -103.4053
7.5833 | -103.4189
7.5871 | -103.4343
7.5908 | -103.4457
7.5946 | -103.4573
7.5984 | -103.4694
7.6022 | -103.4788
7.6059 | -103.4887
7.6097 | -103.4956
7.6135 | -103.4983
7.6173 | -103.5026
7.621 | -103.4905
7.6248 | -103.4805
7.6286 | -103.4579
7.6323 | -103.4356
7.6361 | -103.4101
7.6399 | -103.366
7.6437 | -103.3152
7.6474 | -103.2647
7.6512 | -103.2118
7.655 | -103.1521
7.6588 | -103.0924
7.6625 | -103.0351
7.6663 | -102.9814
7.6701 | -102.9304
7.6739 | -102.8812
7.6776 | -102.8373
7.6814 | -102.8055
7.6852 | -102.7785
7.689 | -102.7573
7.6927 | -102.7475
7.6965 | -102.7472
7.7003 | -102.7563
7.7041 | -102.7756
7.7078 | -102.802
7.7116 | -102.8363
7.7154 | -102.8761
7.7192 | -102.9216
7.7229 | -102.9729
7.7267 | -103.0266
7.7305 | -103.0815
7.7343 | -103.1407
7.738 | -103.1976
7.7418 | -103.2603
7.7456 | -103.3254
7.7494 | -103.3807
7.7531 | -103.4351
7.7569 | -103.4933
7.7607 | -103.5361
7.7645 | -103.5666
7.7682 | -103.5952
7.772 | -103.6188
7.7758 | -103.6355
7.7796 | -103.6448
7.7833 | -103.6501
7.7871 | -103.6532
7.7909 | -103.654
7.7947 | -103.6525
7.7984 | -103.6443
7.8022 | -103.6366
7.806 | -103.6307
7.8098 | -103.6227
7.8135 | -103.6116
7.8173 | -103.6001
7.8211 | -103.5869
7.8249 | -103.572
7.8286 | -103.5563
7.8324 | -103.5397
7.8362 | -103.5149
7.84 | -103.493
7.8437 | -103.4684
7.8475 | -103.4511
7.8513 | -103.4357
7.8551 | -103.4191
7.8588 | -103.4049
7.8626 | -103.402
7.8664 | -103.4039
7.8702 | -103.4016
7.8739 | -103.4087
7.8777 | -103.426
7.8815 | -103.4419
7.8853 | -103.4575
7.889 | -103.4759
7.8928 | -103.4947
7.8966 | -103.5159
7.9003 | -103.5353
7.9041 | -103.556
7.9079 | -103.5792
7.9117 | -103.6034
7.9154 | -103.624
7.9192 | -103.6448
7.923 | -103.6642
7.9268 | -103.6797
7.9305 | -103.6917
7.9343 | -103.7041
7.9381 | -103.7164
7.9419 | -103.7229
7.9456 | -103.72
7.9494 | -103.7167
7.9532 | -103.7151
7.957 | -103.7034
7.9607 | -103.6909
7.9645 | -103.6722
7.9683 | -103.6543
7.9721 | -103.6426
7.9758 | -103.6334
7.9796 | -103.6228
7.9834 | -103.618
7.9872 | -103.6205
7.9909 | -103.6222
7.9947 | -103.629
7.9985 | -103.6404
8.0023 | -103.6499
8.006 | -103.6563
8.0098 | -103.6629
8.0136 | -103.6663
8.0174 | -103.6661
8.0211 | -103.6622
8.0249 | -103.6543
8.0287 | -103.645
8.0325 | -103.635
8.0362 | -103.6272
8.04 | -103.6157
8.0438 | -103.6071
8.0476 | -103.605
8.0513 | -103.6024
8.0551 | -103.6016
8.0589 | -103.6022
8.0627 | -103.6138
8.0664 | -103.6249
8.0702 | -103.6393
8.074 | -103.6614
8.0778 | -103.6856
8.0815 | -103.7114
8.0853 | -103.7421
8.0891 | -103.7736
8.0929 | -103.8069
8.0966 | -103.8331
8.1004 | -103.8571
8.1042 | -103.8756
8.108 | -103.8904
8.1117 | -103.9051
8.1155 | -103.9201
8.1193 | -103.9359
8.1231 | -103.9458
8.1268 | -103.9485
8.1306 | -103.9482
8.1344 | -103.9448
8.1382 | -103.9349
8.1419 | -103.9172
8.1457 | -103.8985
8.1495 | -103.8748
8.1533 | -103.8482
8.157 | -103.8268
8.1608 | -103.8052
8.1646 | -103.7822
8.1684 | -103.7608
8.1721 | -103.734
8.1759 | -103.7137
8.1797 | -103.7015
8.1834 | -103.683
8.1872 | -103.6635
8.191 | -103.6449
8.1948 | -103.6249
8.1985 | -103.6071
8.2023 | -103.5939
8.2061 | -103.5841
8.2099 | -103.576
8.2136 | -103.575
8.2174 | -103.5855
8.2212 | -103.5993
8.225 | -103.6195
8.2287 | -103.6429
8.2325 | -103.6667
8.2363 | -103.6925
8.2401 | -103.7255
8.2438 | -103.7591
8.2476 | -103.794
8.2514 | -103.8223
8.2552 | -103.8467
8.2589 | -103.8652
8.2627 | -103.8781
8.2665 | -103.8804
8.2703 | -103.8797
8.274 | -103.8734
8.2778 | -103.8593
8.2816 | -103.8441
8.2854 | -103.8302
8.2891 | -103.8182
8.2929 | -103.8116
8.2967 | -103.8004
8.3005 | -103.7853
8.3042 | -103.7742
8.308 | -103.7702
8.3118 | -103.7686
8.3156 | -103.769
8.3193 | -103.767
8.3231 | -103.7655
8.3269 | -103.7633
8.3307 | -103.7572
8.3344 | -103.7489
8.3382 | -103.7316
8.342 | -103.705
8.3458 | -103.6794
8.3495 | -103.6531
8.3533 | -103.6243
8.3571 | -103.5952
8.3609 | -103.5742
8.3646 | -103.5532
8.3684 | -103.5391
8.3722 | -103.5302
8.376 | -103.5279
8.3797 | -103.5293
8.3835 | -103.5395
8.3873 | -103.557
8.3911 | -103.5809
8.3948 | -103.6063
8.3986 | -103.6326
8.4024 | -103.6616
8.4062 | -103.6878
8.4099 | -103.7129
8.4137 | -103.7369
8.4175 | -103.7586
8.4213 | -103.7789
8.425 | -103.7905
8.4288 | -103.7929
8.4326 | -103.7859
8.4363 | -103.7785
8.4401 | -103.7653
8.4439 | -103.7535
8.4477 | -103.7363
8.4514 | -103.7193
8.4552 | -103.7071
8.459 | -103.7005
8.4628 | -103.7003
8.4665 | -103.7095
8.4703 | -103.7235
8.4741 | -103.7429
8.4779 | -103.7716
8.4816 | -103.8118
8.4854 | -103.855
8.4892 | -103.9084
8.493 | -103.969
8.4967 | -104.0278
8.5005 | -104.0941
8.5043 | -104.16
8.5081 | -104.2192
8.5118 | -104.2737
8.5156 | -104.3174
8.5194 | -104.3548
8.5232 | -104.3881
8.5269 | -104.4117
8.5307 | -104.4251
8.5345 | -104.4278
8.5383 | -104.4198
8.542 | -104.4063
8.5458 | -104.395
8.5496 | -104.3799
8.5534 | -104.3562
8.5571 | -104.3356
8.5609 | -104.3122
8.5647 | -104.2901
8.5685 | -104.2677
8.5722 | -104.2469
8.576 | -104.2269
8.5798 | -104.2054
8.5836 | -104.183
8.5873 | -104.1599
8.5911 | -104.1317
8.5949 | -104.1022
8.5987 | -104.0726
8.6024 | -104.0427
8.6062 | -104.0056
8.61 | -103.9732
8.6138 | -103.9416
8.6175 | -103.9089
8.6213 | -103.8824
8.6251 | -103.8523
8.6289 | -103.8252
8.6326 | -103.8056
8.6364 | -103.7903
8.6402 | -103.7797
8.644 | -103.7716
8.6477 | -103.767
8.6515 | -103.772
8.6553 | -103.7862
8.6591 | -103.8034
8.6628 | -103.8258
8.6666 | -103.8514
8.6704 | -103.8812
8.6742 | -103.9168
8.6779 | -103.9564
8.6817 | -103.9927
8.6855 | -104.0305
8.6893 | -104.0649
8.693 | -104.0996
8.6968 | -104.1316
8.7006 | -104.1574
8.7043 | -104.1795
8.7081 | -104.1957
8.7119 | -104.1909
8.7157 | -104.2008
8.7194 | -104.2129
8.7232 | -104.2208
8.727 | -104.2243
8.7308 | -104.2245
8.7345 | -104.2265
8.7383 | -104.2279
8.7421 | -104.2264
8.7459 | -104.2234
8.7496 | -104.2193
8.7534 | -104.2211
8.7572 | -104.2246
8.761 | -104.2248
8.7647 | -104.2262
8.7685 | -104.226
8.7723 | -104.2269
8.7761 | -104.2281
8.7798 | -104.2296
8.7836 | -104.2313
8.7874 | -104.2383
8.7912 | -104.2485
8.7949 | -104.2584
8.7987 | -104.263
8.8025 | -104.2687
8.8063 | -104.2783
8.81 | -104.2903
8.8138 | -104.2969
8.8176 | -104.303
8.8214 | -104.3057
8.8251 | -104.3058
8.8289 | -104.3028
8.8327 | -104.2948
8.8365 | -104.282
8.8402 | -104.2705
8.844 | -104.2526
8.8478 | -104.2336
8.8516 | -104.2102
8.8553 | -104.1917
8.8591 | -104.1757
8.8629 | -104.1661
8.8667 | -104.1621
8.8704 | -104.1573
8.8742 | -104.1569
8.878 | -104.1632
8.8818 | -104.1683
8.8855 | -104.1694
8.8893 | -104.1747
8.8931 | -104.1837
8.8969 | -104.1938
8.9006 | -104.2101
8.9044 | -104.2294
8.9082 | -104.2514
8.912 | -104.2733
8.9157 | -104.3004
8.9195 | -104.3297
8.9233 | -104.3598
8.9271 | -104.397
8.9308 | -104.4354
8.9346 | -104.4732
8.9384 | -104.5075
8.9422 | -104.5384
8.9459 | -104.5674
8.9497 | -104.5928
8.9535 | -104.6114
8.9573 | -104.632
8.961 | -104.6463
8.9648 | -104.651
8.9686 | -104.6474
8.9724 | -104.6322
8.9761 | -104.6173
8.9799 | -104.5927
8.9837 | -104.5646
8.9874 | -104.5329
8.9912 | -104.4977
8.995 | -104.4624
8.9988 | -104.4204
9.0025 | -104.382
9.0063 | -104.3514
9.0101 | -104.3197
9.0139 | -104.2895
9.0176 | -104.2657
9.0214 | -104.2436
9.0252 | -104.2254
9.029 | -104.2124
9.0327 | -104.2051
9.0365 | -104.2014
9.0403 | -104.2085
9.0441 | -104.2238
9.0478 | -104.2437
9.0516 | -104.2649
9.0554 | -104.2952
9.0592 | -104.3251
9.0629 | -104.3603
9.0667 | -104.3931
9.0705 | -104.4217
9.0743 | -104.4455
9.078 | -104.466
9.0818 | -104.4774
9.0856 | -104.4834
9.0894 | -104.4755
9.0931 | -104.4611
9.0969 | -104.4443
9.1007 | -104.4213
9.1045 | -104.3911
9.1082 | -104.3526
9.112 | -104.3066
9.1158 | -104.261
9.1196 | -104.2123
9.1233 | -104.1667
9.1271 | -104.1221
9.1309 | -104.0786
9.1347 | -104.0369
9.1384 | -103.9972
9.1422 | -103.9652
9.146 | -103.9468
9.1498 | -103.9319
9.1535 | -103.9224
9.1573 | -103.924
9.1611 | -103.9333
9.1649 | -103.9454
9.1686 | -103.9648
9.1724 | -103.9853
9.1762 | -104.015
9.18 | -104.0521
9.1837 | -104.0928
9.1875 | -104.128
9.1913 | -104.1626
9.1951 | -104.1944
9.1988 | -104.224
9.2026 | -104.2523
9.2064 | -104.2776
9.2102 | -104.2932
9.2139 | -104.3072
9.2177 | -104.3178
9.2215 | -104.3265
9.2253 | -104.3317
9.229 | -104.3283
9.2328 | -104.3241
9.2366 | -104.3205
9.2404 | -104.3188
9.2441 | -104.3224
9.2479 | -104.325
9.2517 | -104.3293
9.2554 | -104.3427
9.2592 | -104.3462
9.263 | -104.3484
9.2668 | -104.3571
9.2705 | -104.3644
9.2743 | -104.3663
9.2781 | -104.3658
9.2819 | -104.3596
9.2856 | -104.3526
9.2894 | -104.3459
9.2932 | -104.3353
9.297 | -104.3235
9.3007 | -104.3123
9.3045 | -104.2932
9.3083 | -104.2735
9.3121 | -104.2485
9.3158 | -104.2242
9.3196 | -104.2036
9.3234 | -104.1867
9.3272 | -104.1719
9.3309 | -104.1598
9.3347 | -104.1459
9.3385 | -104.1319
9.3423 | -104.1182
9.346 | -104.1118
9.3498 | -104.1079
9.3536 | -104.1071
9.3574 | -104.1099
9.3611 | -104.1123
9.3649 | -104.1196
9.3687 | -104.1291
9.3725 | -104.1375
9.3762 | -104.1492
9.38 | -104.1619
9.3838 | -104.1719
9.3876 | -104.184
9.3913 | -104.1985
9.3951 | -104.2138
9.3989 | -104.2281
9.4027 | -104.2412
9.4064 | -104.2498
9.4102 | -104.2561
9.414 | -104.2625
9.4178 | -104.2687
9.4215 | -104.2677
9.4253 | -104.2623
9.4291 | -104.2556
9.4329 | -104.2484
9.4366 | -104.2357
9.4404 | -104.2255
9.4442 | -104.2201
9.448 | -104.2146
9.4517 | -104.2082
9.4555 | -104.2028
9.4593 | -104.1993
9.4631 | -104.1923
9.4668 | -104.1873
9.4706 | -104.1848
9.4744 | -104.1862
9.4782 | -104.1904
9.4819 | -104.1958
9.4857 | -104.1979
9.4895 | -104.1967
9.4933 | -104.194
9.497 | -104.1919
9.5008 | -104.1932
9.5046 | -104.1901
9.5084 | -104.1726
9.5121 | -104.1535
9.5159 | -104.13
9.5197 | -104.1062
9.5235 | -104.0789
9.5272 | -104.0507
9.531 | -104.0211
9.5348 | -103.9943
9.5385 | -103.9721
9.5423 | -103.9533
9.5461 | -103.9399
9.5499 | -103.9326
9.5536 | -103.929
9.5574 | -103.933
9.5612 | -103.9489
9.565 | -103.9744
9.5687 | -104.0044
9.5725 | -104.0389
9.5763 | -104.0828
9.5801 | -104.1304
9.5838 | -104.1816
9.5876 | -104.2298
9.5914 | -104.2778
9.5952 | -104.3207
9.5989 | -104.3628
9.6027 | -104.4026
9.6065 | -104.4429
9.6103 | -104.4772
9.614 | -104.5088
9.6178 | -104.5311
9.6216 | -104.5449
9.6254 | -104.5541
9.6291 | -104.5611
9.6329 | -104.5641
9.6367 | -104.5619
9.6405 | -104.5554
9.6442 | -104.5526
9.648 | -104.5539
9.6518 | -104.5435
9.6556 | -104.5285
9.6593 | -104.5174
9.6631 | -104.5099
9.6669 | -104.4958
9.6707 | -104.4811
9.6744 | -104.466
9.6782 | -104.4518
9.682 | -104.4369
9.6858 | -104.4221
9.6895 | -104.4065
9.6933 | -104.3935
9.6971 | -104.3752
9.7009 | -104.3613
9.7046 | -104.3519
9.7084 | -104.3422
9.7122 | -104.3353
9.716 | -104.3351
9.7197 | -104.3392
9.7235 | -104.3463
9.7273 | -104.3608
9.7311 | -104.3792
9.7348 | -104.404
9.7386 | -104.4277
9.7424 | -104.4544
9.7462 | -104.4767
9.7499 | -104.4956
9.7537 | -104.5153
9.7575 | -104.5405
9.7613 | -104.5668
9.765 | -104.5914
9.7688 | -104.6127
9.7726 | -104.6342
9.7764 | -104.6577
9.7801 | -104.676
9.7839 | -104.6911
9.7877 | -104.7104
9.7914 | -104.7308
9.7952 | -104.7496
9.799 | -104.7653
9.8028 | -104.7765
9.8065 | -104.7863
9.8103 | -104.8009
9.8141 | -104.816
9.8179 | -104.8311
9.8216 | -104.8426
9.8254 | -104.8502
9.8292 | -104.8566
9.833 | -104.8632
9.8367 | -104.8644
9.8405 | -104.8654
9.8443 | -104.8582
9.8481 | -104.8468
9.8518 | -104.8371
9.8556 | -104.8295
9.8594 | -104.8178
9.8632 | -104.8055
9.8669 | -104.7955
9.8707 | -104.7871
9.8745 | -104.7753
9.8783 | -104.7607
9.882 | -104.7438
9.8858 | -104.7298
9.8896 | -104.7144
9.8934 | -104.7047
9.8971 | -104.6949
9.9009 | -104.6796
9.9047 | -104.6666
9.9085 | -104.653
9.9122 | -104.6401
9.916 | -104.6256
9.9198 | -104.6108
9.9236 | -104.6018
9.9273 | -104.5955
9.9311 | -104.5858
9.9349 | -104.584
9.9387 | -104.5838
9.9424 | -104.5855
9.9462 | -104.5922
9.95 | -104.6028
9.9538 | -104.6126
9.9575 | -104.6196
9.9613 | -104.6255
9.9651 | -104.6281
9.9689 | -104.6354
9.9726 | -104.6461
9.9764 | -104.6522
9.9802 | -104.652
9.984 | -104.6523
9.9877 | -104.65
9.9915 | -104.6498
9.9953 | -104.6473
9.9991 | -104.6469
10.0028 | -104.6421
10.0066 | -104.6436
10.0104 | -104.6491
10.0142 | -104.6546
10.0179 | -104.6633
10.0217 | -104.6803
10.0255 | -104.701
10.0293 | -104.7221
10.033 | -104.7426
10.0368 | -104.7696
10.0406 | -104.7943
10.0444 | -104.8192
10.0481 | -104.8457
10.0519 | -104.8712
10.0557 | -104.8965
10.0595 | -104.9137
10.0632 | -104.9292
10.067 | -104.9415
10.0708 | -104.9498
10.0745 | -104.9504
10.0783 | -104.9498
10.0821 | -104.9416
10.0859 | -104.9282
10.0896 | -104.9134
10.0934 | -104.8975
10.0972 | -104.8788
10.101 | -104.864
10.1047 | -104.8536
10.1085 | -104.8453
10.1123 | -104.8436
10.1161 | -104.8399
10.1198 | -104.8368
10.1236 | -104.8403
10.1274 | -104.8527
10.1312 | -104.8666
10.1349 | -104.8809
10.1387 | -104.8937
10.1425 | -104.9126
10.1463 | -104.932
10.15 | -104.9523
10.1538 | -104.968
10.1576 | -104.9815
10.1614 | -104.9907
10.1651 | -104.9935
10.1689 | -104.9936
10.1727 | -104.9902
10.1765 | -104.986
10.1802 | -104.9802
10.184 | -104.9684
10.1878 | -104.9535
10.1916 | -104.9343
10.1953 | -104.914
10.1991 | -104.891
10.2029 | -104.8687
10.2067 | -104.8501
10.2104 | -104.8278
10.2142 | -104.8107
10.218 | -104.7907
10.2218 | -104.7776
10.2255 | -104.7679
10.2293 | -104.7598
10.2331 | -104.7593
10.2369 | -104.7639
10.2406 | -104.7707
10.2444 | -104.7837
10.2482 | -104.796
10.252 | -104.8114
10.2557 | -104.8306
10.2595 | -104.854
10.2633 | -104.8779
10.2671 | -104.9008
10.2708 | -104.9221
10.2746 | -104.9407
10.2784 | -104.9571
10.2822 | -104.9714
10.2859 | -104.9846
10.2897 | -104.9947
10.2935 | -104.9996
10.2973 | -105.0049
10.301 | -105.0096
10.3048 | -105.015
10.3086 | -105.0243
10.3124 | -105.0355
10.3161 | -105.0436
10.3199 | -105.0481
10.3237 | -105.0535
10.3275 | -105.0574
10.3312 | -105.0656
10.335 | -105.0778
10.3388 | -105.0911
10.3425 | -105.1062
10.3463 | -105.1214
10.3501 | -105.1386
10.3539 | -105.1514
10.3576 | -105.1594
10.3614 | -105.1659
10.3652 | -105.168
10.369 | -105.1727
10.3727 | -105.1698
10.3765 | -105.159
10.3803 | -105.1456
10.3841 | -105.1269
10.3878 | -105.1012
10.3916 | -105.0737
10.3954 | -105.0462
10.3992 | -105.0177
10.4029 | -104.9835
10.4067 | -104.9461
10.4105 | -104.9048
10.4143 | -104.8661
10.418 | -104.8334
10.4218 | -104.8019
10.4256 | -104.7699
10.4294 | -104.745
10.4331 | -104.7213
10.4369 | -104.7017
10.4407 | -104.6861
10.4445 | -104.6764
10.4482 | -104.6747
10.452 | -104.6763
10.4558 | -104.6804
10.4596 | -104.6875
10.4633 | -104.6941
10.4671 | -104.7045
10.4709 | -104.7207
10.4747 | -104.7345
10.4784 | -104.7482
10.4822 | -104.7627
10.486 | -104.774
10.4898 | -104.7844
10.4935 | -104.7951
10.4973 | -104.8028
10.5011 | -104.8062
10.5049 | -104.8046
10.5086 | -104.8043
10.5124 | -104.8047
10.5162 | -104.8041
10.52 | -104.7974
10.5237 | -104.788
10.5275 | -104.776
10.5313 | -104.7659
10.5351 | -104.7538
10.5388 | -104.7423
10.5426 | -104.7305
10.5464 | -104.7211
10.5502 | -104.7095
10.5539 | -104.7003
10.5577 | -104.6927
10.5615 | -104.6868
10.5653 | -104.6844
10.569 | -104.6845
10.5728 | -104.6831
10.5766 | -104.684
10.5804 | -104.6863
10.5841 | -104.6879
10.5879 | -104.6924
10.5917 | -104.6959
10.5955 | -104.6999
10.5992 | -104.7046
10.603 | -104.7031
10.6068 | -104.705
10.6105 | -104.7091
10.6143 | -104.7139
10.6181 | -104.7229
10.6219 | -104.7283
10.6256 | -104.7263
10.6294 | -104.73
10.6332 | -104.7357
10.637 | -104.7365
10.6407 | -104.7394
10.6445 | -104.7411
10.6483 | -104.7417
10.6521 | -104.7433
10.6558 | -104.7497
10.6596 | -104.7571
10.6634 | -104.7631
10.6672 | -104.7711
10.6709 | -104.7771
10.6747 | -104.784
10.6785 | -104.7927
10.6823 | -104.8059
10.686 | -104.818
10.6898 | -104.8302
10.6936 | -104.8414
10.6974 | -104.8508
10.7011 | -104.8518
10.7049 | -104.851
10.7087 | -104.8513
10.7125 | -104.8496
10.7162 | -104.8486
10.72 | -104.8455
10.7238 | -104.8373
10.7276 | -104.8282
10.7313 | -104.8187
10.7351 | -104.8107
10.7389 | -104.8052
10.7427 | -104.7976
10.7464 | -104.7871
10.7502 | -104.7761
10.754 | -104.7649
10.7578 | -104.7588
10.7615 | -104.7481
10.7653 | -104.7353
10.7691 | -104.7234
10.7729 | -104.7124
10.7766 | -104.7069
10.7804 | -104.7047
10.7842 | -104.6977
10.788 | -104.6914
10.7917 | -104.6873
10.7955 | -104.6856
10.7993 | -104.6861
10.8031 | -104.6843
10.8068 | -104.6821
10.8106 | -104.6795
10.8144 | -104.679
10.8182 | -104.6822
10.8219 | -104.6844
10.8257 | -104.6897
10.8295 | -104.695
10.8333 | -104.6999
10.837 | -104.7037
10.8408 | -104.7125
10.8446 | -104.7217
10.8484 | -104.7313
10.8521 | -104.747
10.8559 | -104.7716
10.8597 | -104.7973
10.8635 | -104.8202
10.8672 | -104.8439
10.871 | -104.866
10.8748 | -104.8863
10.8786 | -104.9088
10.8823 | -104.9316
10.8861 | -104.9548
10.8899 | -104.9808
10.8936 | -105.0042
10.8974 | -105.0246
10.9012 | -105.0464
10.905 | -105.0699
10.9087 | -105.0909
10.9125 | -105.1139
10.9163 | -105.1343
10.9201 | -105.1587
10.9238 | -105.1824
10.9276 | -105.2048
10.9314 | -105.2236
10.9352 | -105.2393
10.9389 | -105.257
10.9427 | -105.2699
10.9465 | -105.2798
10.9503 | -105.2865
10.954 | -105.2898
10.9578 | -105.295
10.9616 | -105.3021
10.9654 | -105.306
10.9691 | -105.3054
10.9729 | -105.303
10.9767 | -105.3017
10.9805 | -105.2954
10.9842 | -105.2851
10.988 | -105.2795
10.9918 | -105.2723
10.9956 | -105.2644
10.9993 | -105.2498
11.0031 | -105.2348
11.0069 | -105.2188
11.0107 | -105.1998
11.0144 | -105.1784
11.0182 | -105.1565
11.022 | -105.134
11.0258 | -105.1108
11.0295 | -105.0862
11.0333 | -105.0657
11.0371 | -105.0429
11.0409 | -105.0208
11.0446 | -104.9982
11.0484 | -104.9754
11.0522 | -104.9502
11.056 | -104.9313
11.0597 | -104.9158
11.0635 | -104.9027
11.0673 | -104.8923
11.0711 | -104.8834
11.0748 | -104.8767
11.0786 | -104.872
11.0824 | -104.8687
11.0862 | -104.8676
11.0899 | -104.8677
11.0937 | -104.8718
11.0975 | -104.8733
11.1013 | -104.8753
11.105 | -104.8839
11.1088 | -104.8914
11.1126 | -104.9012
11.1164 | -104.9076
11.1201 | -104.9082
11.1239 | -104.9042
11.1277 | -104.9068
11.1315 | -104.9115
11.1352 | -104.9146
11.139 | -104.914
11.1428 | -104.9148
11.1466 | -104.9178
11.1503 | -104.9172
11.1541 | -104.9175
11.1579 | -104.9197
11.1616 | -104.9225
11.1654 | -104.9265
11.1692 | -104.9346
11.173 | -104.9471
11.1767 | -104.9587
11.1805 | -104.9709
11.1843 | -104.9821
11.1881 | -104.9986
11.1918 | -105.0163
11.1956 | -105.0274
11.1994 | -105.0415
11.2032 | -105.0551
11.2069 | -105.0643
11.2107 | -105.0699
11.2145 | -105.0783
11.2183 | -105.0768
11.222 | -105.0731
11.2258 | -105.0646
11.2296 | -105.0541
11.2334 | -105.0418
11.2371 | -105.0215
11.2409 | -104.9956
11.2447 | -104.9683
11.2485 | -104.9458
11.2522 | -104.9224
11.256 | -104.9027
11.2598 | -104.8872
11.2636 | -104.8736
11.2673 | -104.8586
11.2711 | -104.845
11.2749 | -104.8342
11.2787 | -104.8283
11.2824 | -104.8273
11.2862 | -104.8322
11.29 | -104.8338
11.2938 | -104.8352
11.2975 | -104.837
11.3013 | -104.8409
11.3051 | -104.8484
11.3089 | -104.8557
11.3126 | -104.863
11.3164 | -104.8686
11.3202 | -104.8715
11.324 | -104.878
11.3277 | -104.8817
11.3315 | -104.8844
11.3353 | -104.8836
11.3391 | -104.8847
11.3428 | -104.8894
11.3466 | -104.892
11.3504 | -104.8866
11.3542 | -104.8841
11.3579 | -104.8843
11.3617 | -104.8837
11.3655 | -104.8792
11.3693 | -104.8769
11.373 | -104.8747
11.3768 | -104.8721
11.3806 | -104.8732
11.3844 | -104.8699
11.3881 | -104.8694
11.3919 | -104.8695
11.3957 | -104.8668
11.3995 | -104.8709
11.4032 | -104.8788
11.407 | -104.8819
11.4108 | -104.8792
11.4146 | -104.88
11.4183 | -104.8782
11.4221 | -104.8801
11.4259 | -104.8801
11.4297 | -104.8823
11.4334 | -104.8897
11.4372 | -104.897
11.441 | -104.9044
11.4447 | -104.9141
11.4485 | -104.9256
11.4523 | -104.9358
11.4561 | -104.9511
11.4598 | -104.9697
11.4636 | -104.9902
11.4674 | -105.0123
11.4712 | -105.0344
11.4749 | -105.0576
11.4787 | -105.0797
11.4825 | -105.1045
11.4863 | -105.1324
11.49 | -105.1584
11.4938 | -105.1841
11.4976 | -105.209
11.5014 | -105.2338
11.5051 | -105.2562
11.5089 | -105.2745
11.5127 | -105.295
11.5165 | -105.3151
11.5202 | -105.3317
11.524 | -105.3473
11.5278 | -105.3527
11.5316 | -105.3549
11.5353 | -105.3558
11.5391 | -105.3542
11.5429 | -105.3496
11.5467 | -105.3415
11.5504 | -105.3359
11.5542 | -105.3262
11.558 | -105.3103
11.5618 | -105.2966
11.5655 | -105.2771
11.5693 | -105.2581
11.5731 | -105.2424
11.5769 | -105.2308
11.5806 | -105.2157
11.5844 | -105.2014
11.5882 | -105.187
11.592 | -105.1713
11.5957 | -105.1565
11.5995 | -105.1416
11.6033 | -105.128
11.6071 | -105.1155
11.6108 | -105.1032
11.6146 | -105.0884
11.6184 | -105.072
11.6222 | -105.06
11.6259 | -105.047
11.6297 | -105.0312
11.6335 | -105.0185
11.6373 | -105.0061
11.641 | -104.9942
11.6448 | -104.9776
11.6486 | -104.9612
11.6524 | -104.9473
11.6561 | -104.931
11.6599 | -104.917
11.6637 | -104.9058
11.6675 | -104.8939
11.6712 | -104.8818
11.675 | -104.8683
11.6788 | -104.8573
11.6826 | -104.8473
11.6863 | -104.8367
11.6901 | -104.8254
11.6939 | -104.8141
11.6976 | -104.8066
11.7014 | -104.7996
11.7052 | -104.7908
11.709 | -104.784
11.7127 | -104.7752
11.7165 | -104.7701
11.7203 | -104.7655
11.7241 | -104.7617
11.7278 | -104.7634
11.7316 | -104.7632
11.7354 | -104.7654
11.7392 | -104.7677
11.7429 | -104.7731
11.7467 | -104.7766
11.7505 | -104.7818
11.7543 | -104.7871
11.758 | -104.7953
11.7618 | -104.8057
11.7656 | -104.8103
11.7694 | -104.8147
11.7731 | -104.8161
11.7769 | -104.8176
11.7807 | -104.8207
11.7845 | -104.8244
11.7882 | -104.8253
11.792 | -104.8266
11.7958 | -104.8301
11.7996 | -104.8375
11.8033 | -104.8401
11.8071 | -104.8424
11.8109 | -104.8468
11.8147 | -104.8577
11.8184 | -104.8706
11.8222 | -104.8809
11.826 | -104.8923
11.8298 | -104.906
11.8335 | -104.9211
11.8373 | -104.9355
11.8411 | -104.9519
11.8449 | -104.9707
11.8486 | -104.988
11.8524 | -105.0023
11.8562 | -105.0179
11.86 | -105.031
11.8637 | -105.0453
11.8675 | -105.0608
11.8713 | -105.0742
11.8751 | -105.0871
11.8788 | -105.0991
11.8826 | -105.1139
11.8864 | -105.1282
11.8902 | -105.142
11.8939 | -105.1596
11.8977 | -105.1785
11.9015 | -105.1988
11.9053 | -105.2152
11.909 | -105.2322
11.9128 | -105.2468
11.9166 | -105.264
11.9204 | -105.2819
11.9241 | -105.2992
11.9279 | -105.3144
11.9317 | -105.3295
11.9355 | -105.3465
11.9392 | -105.3593
11.943 | -105.3691
11.9468 | -105.3752
11.9506 | -105.3817
11.9543 | -105.3894
11.9581 | -105.3972
11.9619 | -105.3996
11.9657 | -105.3972
11.9694 | -105.3956
11.9732 | -105.3984
11.977 | -105.3907
11.9807 | -105.3925
11.9845 | -105.391
11.9883 | -105.3836
11.9921 | -105.3716
11.9958 | -105.3606
11.9996 | -105.351
12.0034 | -105.3334
12.0072 | -105.3148
12.0109 | -105.2974
12.0147 | -105.2771
12.0185 | -105.2562
12.0223 | -105.2335
12.026 | -105.2119
12.0298 | -105.1876
12.0336 | -105.163
12.0374 | -105.1426
12.0411 | -105.1212
12.0449 | -105.1
12.0487 | -105.0785
12.0525 | -105.0552
12.0562 | -105.0331
12.06 | -105.0079
12.0638 | -104.9817
12.0676 | -104.959
12.0713 | -104.9345
12.0751 | -104.9148
12.0789 | -104.9021
12.0827 | -104.8914
12.0864 | -104.8807
12.0902 | -104.8696
12.094 | -104.8624
12.0978 | -104.8607
12.1015 | -104.8608
12.1053 | -104.8651
12.1091 | -104.8763
12.1129 | -104.8913
12.1166 | -104.9119
12.1204 | -104.9358
12.1242 | -104.9589
12.128 | -104.9838
12.1317 | -105.0139
12.1355 | -105.0476
12.1393 | -105.0829
12.1431 | -105.1192
12.1468 | -105.1588
12.1506 | -105.2
12.1544 | -105.2391
12.1582 | -105.279
12.1619 | -105.3181
12.1657 | -105.3578
12.1695 | -105.3974
12.1733 | -105.4348
12.177 | -105.467
12.1808 | -105.498
12.1846 | -105.5211
12.1884 | -105.5451
12.1921 | -105.5658
12.1959 | -105.578
12.1997 | -105.5924
12.2035 | -105.6009
12.2072 | -105.6071
12.211 | -105.6145
12.2148 | -105.6103
12.2186 | -105.6042
12.2223 | -105.5985
12.2261 | -105.5927
12.2299 | -105.5871
12.2337 | -105.5783
12.2374 | -105.5725
12.2412 | -105.5683
12.245 | -105.5583
12.2487 | -105.5462
12.2525 | -105.5355
12.2563 | -105.5198
12.2601 | -105.5043
12.2638 | -105.4941
12.2676 | -105.4825
12.2714 | -105.472
12.2752 | -105.4603
12.2789 | -105.449
12.2827 | -105.4384
12.2865 | -105.4252
12.2903 | -105.4131
12.294 | -105.4048
12.2978 | -105.3972
12.3016 | -105.3899
12.3054 | -105.3856
12.3091 | -105.3831
12.3129 | -105.3776
12.3167 | -105.3743
12.3205 | -105.3713
12.3242 | -105.3669
12.328 | -105.3628
12.3318 | -105.3584
12.3356 | -105.3558
12.3393 | -105.3518
12.3431 | -105.3501
12.3469 | -105.3481
12.3507 | -105.3427
12.3544 | -105.3402
12.3582 | -105.3401
12.362 | -105.3402
12.3658 | -105.3344
12.3695 | -105.3308
12.3733 | -105.3305
12.3771 | -105.3251
12.3809 | -105.3209
12.3846 | -105.3193
12.3884 | -105.314
12.3922 | -105.3119
12.396 | -105.3099
12.3997 | -105.3075
12.4035 | -105.307
12.4073 | -105.3077
12.4111 | -105.3096
12.4148 | -105.3121
12.4186 | -105.3177
12.4224 | -105.322
12.4262 | -105.3271
12.4299 | -105.3301
12.4337 | -105.3347
12.4375 | -105.3384
12.4413 | -105.3406
12.445 | -105.3462
12.4488 | -105.3518
12.4526 | -105.3569
12.4564 | -105.3596
12.4601 | -105.3579
12.4639 | -105.3544
12.4677 | -105.3494
12.4715 | -105.3414
12.4752 | -105.3338
12.479 | -105.3275
12.4828 | -105.316
12.4866 | -105.303
12.4903 | -105.293
12.4941 | -105.2822
12.4979 | -105.2686
12.5017 | -105.2526
12.5054 | -105.2367
12.5092 | -105.2232
12.513 | -105.2075
12.5168 | -105.1961
12.5205 | -105.1856
12.5243 | -105.1812
12.5281 | -105.1745
12.5318 | -105.1702
12.5356 | -105.1655
12.5394 | -105.1618
12.5432 | -105.1607
12.5469 | -105.1618
12.5507 | -105.1637
12.5545 | -105.1667
12.5583 | -105.1719
12.562 | -105.1778
12.5658 | -105.1867
12.5696 | -105.1995
12.5734 | -105.2132
12.5771 | -105.2289
12.5809 | -105.2434
12.5847 | -105.257
12.5885 | -105.2729
12.5922 | -105.2895
12.596 | -105.3064
12.5998 | -105.3213
12.6036 | -105.3372
12.6073 | -105.3547
12.6111 | -105.3725
12.6149 | -105.3907
12.6187 | -105.4054
12.6224 | -105.4199
12.6262 | -105.4364
12.63 | -105.4514
12.6338 | -105.4658
12.6375 | -105.4816
12.6413 | -105.4986
12.6451 | -105.5129
12.6489 | -105.5268
12.6526 | -105.5357
12.6564 | -105.5453
12.6602 | -105.5578
12.664 | -105.5696
12.6677 | -105.5801
12.6715 | -105.5903
12.6753 | -105.6028
12.6791 | -105.6172
12.6828 | -105.6318
12.6866 | -105.6473
12.6904 | -105.6626
12.6942 | -105.6749
12.6979 | -105.6859
12.7017 | -105.6973
12.7055 | -105.7072
12.7093 | -105.7198
12.713 | -105.7311
12.7168 | -105.739
12.7206 | -105.7452
12.7244 | -105.7512
12.7281 | -105.7591
12.7319 | -105.7619
12.7357 | -105.7652
12.7395 | -105.7644
12.7432 | -105.76
12.747 | -105.7554
12.7508 | -105.7465
12.7546 | -105.7358
12.7583 | -105.7257
12.7621 | -105.7101
12.7659 | -105.6969
12.7697 | -105.6815
12.7734 | -105.6626
12.7772 | -105.6427
12.781 | -105.6268
12.7848 | -105.6078
12.7885 | -105.5899
12.7923 | -105.5727
12.7961 | -105.5561
12.7998 | -105.5396
12.8036 | -105.5249
12.8074 | -105.5092
12.8112 | -105.4911
12.8149 | -105.4779
12.8187 | -105.4678
12.8225 | -105.4622
12.8263 | -105.457
12.83 | -105.453
12.8338 | -105.4494
12.8376 | -105.4511
12.8414 | -105.4523
12.8451 | -105.4521
12.8489 | -105.4547
12.8527 | -105.4582
12.8565 | -105.4642
12.8602 | -105.4696
12.864 | -105.4754
12.8678 | -105.4832
12.8716 | -105.4901
12.8753 | -105.4966
12.8791 | -105.5058
12.8829 | -105.5136
12.8867 | -105.5213
12.8904 | -105.533
12.8942 | -105.5428
12.898 | -105.5524
12.9018 | -105.5649
12.9055 | -105.5766
12.9093 | -105.586
12.9131 | -105.5925
12.9169 | -105.6008
12.9206 | -105.6057
12.9244 | -105.6144
12.9282 | -105.6241
12.932 | -105.6346
12.9357 | -105.6438
12.9395 | -105.6508
12.9433 | -105.6566
12.9471 | -105.663
12.9508 | -105.6641
12.9546 | -105.6655
12.9584 | -105.668
12.9622 | -105.6684
12.9659 | -105.6647
12.9697 | -105.6643
12.9735 | -105.6619
12.9773 | -105.6574
12.981 | -105.6514
12.9848 | -105.6447
12.9886 | -105.6367
12.9924 | -105.6247
12.9961 | -105.615
12.9999 | -105.6044
13.0037 | -105.5943
13.0075 | -105.5845
13.0112 | -105.5768
13.015 | -105.5695
13.0188 | -105.5605
13.0226 | -105.5507
13.0263 | -105.5393
13.0301 | -105.5284
13.0339 | -105.5183
13.0377 | -105.5099
13.0414 | -105.5001
13.0452 | -105.4884
13.049 | -105.4768
13.0528 | -105.465
13.0565 | -105.4546
13.0603 | -105.4461
13.0641 | -105.4371
13.0678 | -105.4282
13.0716 | -105.421
13.0754 | -105.4129
13.0792 | -105.4004
13.0829 | -105.3915
13.0867 | -105.3851
13.0905 | -105.3806
13.0943 | -105.3786
13.098 | -105.3774
13.1018 | -105.3776
13.1056 | -105.3787
13.1094 | -105.3791
13.1131 | -105.379
13.1169 | -105.3806
13.1207 | -105.3821
13.1245 | -105.3866
13.1282 | -105.3928
13.132 | -105.4002
13.1358 | -105.4078
13.1396 | -105.4161
13.1433 | -105.4253
13.1471 | -105.4336
13.1509 | -105.4451
13.1547 | -105.4558
13.1584 | -105.467
13.1622 | -105.4765
13.166 | -105.4876
13.1698 | -105.497
13.1735 | -105.5075
13.1773 | -105.5186
13.1811 | -105.5305
13.1849 | -105.5354
13.1886 | -105.5373
13.1924 | -105.541
13.1962 | -105.5419
13.2 | -105.5429
13.2037 | -105.5436
13.2075 | -105.5457
13.2113 | -105.5446
13.2151 | -105.5452
13.2188 | -105.544
13.2226 | -105.5443
13.2264 | -105.5431
13.2302 | -105.5411
13.2339 | -105.5397
13.2377 | -105.5374
13.2415 | -105.5379
13.2453 | -105.5382
13.249 | -105.5385
13.2528 | -105.5393
13.2566 | -105.5434
13.2604 | -105.5421
13.2641 | -105.5422
13.2679 | -105.5425
13.2717 | -105.5486
13.2755 | -105.5537
13.2792 | -105.5573
13.283 | -105.5559
13.2868 | -105.5529
13.2906 | -105.5524
13.2943 | -105.5501
13.2981 | -105.5473
13.3019 | -105.5432
13.3057 | -105.5377
13.3094 | -105.5284
13.3132 | -105.5168
13.317 | -105.5073
13.3208 | -105.4968
13.3245 | -105.4872
13.3283 | -105.4781
13.3321 | -105.4678
13.3359 | -105.4578
13.3396 | -105.4495
13.3434 | -105.4415
13.3472 | -105.4365
13.3509 | -105.4322
13.3547 | -105.4328
13.3585 | -105.4342
13.3623 | -105.4378
13.366 | -105.4428
13.3698 | -105.4506
13.3736 | -105.4628
13.3774 | -105.4798
13.3811 | -105.4956
13.3849 | -105.5097
13.3887 | -105.5238
13.3925 | -105.5383
13.3962 | -105.553
13.4 | -105.5657
13.4038 | -105.5783
13.4076 | -105.5883
13.4113 | -105.6001
13.4151 | -105.6043
13.4189 | -105.6074
13.4227 | -105.6082
13.4264 | -105.604
13.4302 | -105.5984
13.434 | -105.5872
13.4378 | -105.5734
13.4415 | -105.5572
13.4453 | -105.5395
13.4491 | -105.5181
13.4529 | -105.496
13.4566 | -105.4692
13.4604 | -105.4403
13.4642 | -105.4108
13.468 | -105.378
13.4717 | -105.3469
13.4755 | -105.321
13.4793 | -105.2955
13.4831 | -105.2695
13.4868 | -105.2434
13.4906 | -105.222
13.4944 | -105.2036
13.4982 | -105.1894
13.5019 | -105.1765
13.5057 | -105.1687
13.5095 | -105.1619
13.5133 | -105.1572
13.517 | -105.155
13.5208 | -105.1552
13.5246 | -105.1581
13.5284 | -105.1623
13.5321 | -105.166
13.5359 | -105.1721
13.5397 | -105.1769
13.5435 | -105.1857
13.5472 | -105.1934
13.551 | -105.202
13.5548 | -105.2094
13.5586 | -105.2167
13.5623 | -105.2229
13.5661 | -105.2264
13.5699 | -105.2281
13.5737 | -105.2311
13.5774 | -105.2334
13.5812 | -105.2368
13.585 | -105.2383
13.5888 | -105.2384
13.5925 | -105.2375
13.5963 | -105.236
13.6001 | -105.2352
13.6038 | -105.2335
13.6076 | -105.2354
13.6114 | -105.2356
13.6152 | -105.2392
13.6189 | -105.2422
13.6227 | -105.2457
13.6265 | -105.2494
13.6303 | -105.2525
13.634 | -105.255
13.6378 | -105.2573
13.6416 | -105.2594
13.6454 | -105.264
13.6491 | -105.2733
13.6529 | -105.28
13.6567 | -105.2863
13.6605 | -105.292
13.6642 | -105.3006
13.668 | -105.3091
13.6718 | -105.3143
13.6756 | -105.3195
13.6793 | -105.3244
13.6831 | -105.3303
13.6869 | -105.336
13.6907 | -105.3404
13.6944 | -105.3438
13.6982 | -105.3486
13.702 | -105.3521
13.7058 | -105.357
13.7095 | -105.3642
13.7133 | -105.3688
13.7171 | -105.3712
13.7209 | -105.3728
13.7246 | -105.3736
13.7284 | -105.3739
13.7322 | -105.3734
13.736 | -105.3729
13.7397 | -105.3674
13.7435 | -105.3619
13.7473 | -105.3555
13.7511 | -105.3505
13.7548 | -105.348
13.7586 | -105.3472
13.7624 | -105.3448
13.7662 | -105.3428
13.7699 | -105.3392
13.7737 | -105.3382
13.7775 | -105.3386
13.7813 | -105.339
13.785 | -105.3384
13.7888 | -105.3372
13.7926 | -105.3387
13.7964 | -105.3403
13.8001 | -105.3379
13.8039 | -105.3357
13.8077 | -105.3359
13.8115 | -105.3313
13.8152 | -105.3247
13.819 | -105.3216
13.8228 | -105.3196
13.8266 | -105.3145
13.8303 | -105.3084
13.8341 | -105.2994
13.8379 | -105.2893
13.8417 | -105.281
13.8454 | -105.2722
13.8492 | -105.2595
13.853 | -105.2511
13.8568 | -105.2414
13.8605 | -105.2306
13.8643 | -105.2212
13.8681 | -105.2094
13.8719 | -105.2034
13.8756 | -105.1968
13.8794 | -105.191
13.8832 | -105.1854
13.8869 | -105.1794
13.8907 | -105.1745
13.8945 | -105.1715
13.8983 | -105.1702
13.902 | -105.1697
13.9058 | -105.1691
13.9096 | -105.1695
13.9134 | -105.1749
13.9171 | -105.1834
13.9209 | -105.1916
13.9247 | -105.1956
13.9285 | -105.2016
13.9322 | -105.2086
13.936 | -105.2154
13.9398 | -105.2232
13.9436 | -105.2305
13.9473 | -105.2359
13.9511 | -105.2394
13.9549 | -105.2431
13.9587 | -105.2384
13.9624 | -105.2238
13.9662 | -105.2062
13.97 | -105.1952
13.9738 | -105.1856
13.9775 | -105.1754
13.9813 | -105.1637
13.9851 | -105.1488
13.9889 | -105.1358
13.9926 | -105.1185
13.9964 | -105.0984
14.0002 | -105.0773
14.004 | -105.051
14.0077 | -105.0244
14.0115 | -104.9995
14.0153 | -104.9748
14.0191 | -104.9477
14.0228 | -104.9215
14.0266 | -104.8966
14.0304 | -104.8747
14.0342 | -104.8557
14.0379 | -104.8378
14.0417 | -104.8222
14.0455 | -104.8092
14.0493 | -104.8017
14.053 | -104.7974
14.0568 | -104.7996
14.0606 | -104.8068
14.0644 | -104.8183
14.0681 | -104.8364
14.0719 | -104.8563
14.0757 | -104.8835
14.0795 | -104.9151
14.0832 | -104.9519
14.087 | -104.9925
14.0908 | -105.0381
14.0946 | -105.0863
14.0983 | -105.1365
14.1021 | -105.1887
14.1059 | -105.2396
14.1097 | -105.2921
14.1134 | -105.3441
14.1172 | -105.3956
14.121 | -105.4461
14.1248 | -105.4947
14.1285 | -105.5435
14.1323 | -105.5882
14.1361 | -105.6338
14.1399 | -105.678
14.1436 | -105.7197
14.1474 | -105.7589
14.1512 | -105.796
14.1549 | -105.8216
14.1587 | -105.8508
14.1625 | -105.886
14.1663 | -105.93
14.17 | -105.9691
14.1738 | -105.9937
14.1776 | -106.0081
14.1814 | -106.0222
14.1851 | -106.0392
14.1889 | -106.0597
14.1927 | -106.0763
14.1965 | -106.0914
14.2002 | -106.1071
14.204 | -106.123
14.2078 | -106.142
14.2116 | -106.1567
14.2153 | -106.1663
14.2191 | -106.1702
14.2229 | -106.176
14.2267 | -106.1833
14.2304 | -106.1906
14.2342 | -106.193
14.238 | -106.1962
14.2418 | -106.2003
14.2455 | -106.2008
14.2493 | -106.199
14.2531 | -106.1953
14.2569 | -106.191
14.2606 | -106.1848
14.2644 | -106.1765
14.2682 | -106.165
14.272 | -106.1548
14.2757 | -106.1438
14.2795 | -106.1302
14.2833 | -106.113
14.2871 | -106.0953
14.2908 | -106.0782
14.2946 | -106.0609
14.2984 | -106.0409
14.3022 | -106.0247
14.3059 | -106.0109
14.3097 | -105.9975
14.3135 | -105.9812
14.3173 | -105.9646
14.321 | -105.9507
14.3248 | -105.9356
14.3286 | -105.9258
14.3324 | -105.912
14.3361 | -105.8979
14.3399 | -105.8826
14.3437 | -105.8696
14.3475 | -105.8584
14.3512 | -105.8478
14.355 | -105.8393
14.3588 | -105.8323
14.3626 | -105.8257
14.3663 | -105.8224
14.3701 | -105.8181
14.3739 | -105.8143
14.3777 | -105.8147
14.3814 | -105.8168
14.3852 | -105.8197
14.389 | -105.8209
14.3928 | -105.8205
14.3965 | -105.8203
14.4003 | -105.8196
14.4041 | -105.8211
14.4079 | -105.8252
14.4116 | -105.8305
14.4154 | -105.8326
14.4192 | -105.8356
14.423 | -105.8418
14.4267 | -105.8462
14.4305 | -105.8481
14.4343 | -105.8471
14.438 | -105.846
14.4418 | -105.8457
14.4456 | -105.8414
14.4494 | -105.8366
14.4531 | -105.8345
14.4569 | -105.8331
14.4607 | -105.831
14.4645 | -105.8266
14.4682 | -105.8231
14.472 | -105.8184
14.4758 | -105.8142
14.4796 | -105.81
14.4833 | -105.8064
14.4871 | -105.8013
14.4909 | -105.7972
14.4947 | -105.7924
14.4984 | -105.7857
14.5022 | -105.777
14.506 | -105.7695
14.5098 | -105.7604
14.5135 | -105.7516
14.5173 | -105.7438
14.5211 | -105.7378
14.5249 | -105.7316
14.5286 | -105.7225
14.5324 | -105.7182
14.5362 | -105.7121
14.54 | -105.7054
14.5437 | -105.6989
14.5475 | -105.6915
14.5513 | -105.684
14.5551 | -105.6766
14.5588 | -105.6704
14.5626 | -105.6646
14.5664 | -105.6566
14.5702 | -105.6511
14.5739 | -105.6445
14.5777 | -105.6369
14.5815 | -105.63
14.5853 | -105.6217
14.589 | -105.6162
14.5928 | -105.6124
14.5966 | -105.6063
14.6004 | -105.6024
14.6041 | -105.6028
14.6079 | -105.6057
14.6117 | -105.6074
14.6155 | -105.6077
14.6192 | -105.6087
14.623 | -105.6129
14.6268 | -105.6192
14.6306 | -105.626
14.6343 | -105.6345
14.6381 | -105.6426
14.6419 | -105.652
14.6457 | -105.66
14.6494 | -105.6672
14.6532 | -105.6764
14.657 | -105.6832
14.6608 | -105.6822
14.6645 | -105.679
14.6683 | -105.6768
14.6721 | -105.6768
14.6759 | -105.6709
14.6796 | -105.6651
14.6834 | -105.6603
14.6872 | -105.6525
14.691 | -105.6432
14.6947 | -105.6323
14.6985 | -105.6196
14.7023 | -105.6045
14.706 | -105.5911
14.7098 | -105.5783
14.7136 | -105.5673
14.7174 | -105.5559
14.7211 | -105.5452
14.7249 | -105.5325
14.7287 | -105.5216
14.7325 | -105.5139
14.7362 | -105.5064
14.74 | -105.4996
14.7438 | -105.4931
14.7476 | -105.4882
14.7513 | -105.4841
14.7551 | -105.4811
14.7589 | -105.4809
14.7627 | -105.4822
14.7664 | -105.4823
14.7702 | -105.4859
14.774 | -105.4887
14.7778 | -105.4909
14.7815 | -105.4934
14.7853 | -105.4982
14.7891 | -105.5066
14.7929 | -105.5114
14.7966 | -105.5159
14.8004 | -105.5202
14.8042 | -105.5243
14.808 | -105.5298
14.8117 | -105.5309
14.8155 | -105.525
14.8193 | -105.5205
14.8231 | -105.516
14.8268 | -105.5115
14.8306 | -105.5077
14.8344 | -105.5033
14.8382 | -105.4967
14.8419 | -105.4893
14.8457 | -105.4824
14.8495 | -105.4757
14.8533 | -105.4677
14.857 | -105.463
14.8608 | -105.4583
14.8646 | -105.4528
14.8684 | -105.4491
14.8721 | -105.4464
14.8759 | -105.4445
14.8797 | -105.4418
14.8835 | -105.436
14.8872 | -105.4316
14.891 | -105.4274
14.8948 | -105.4174
14.8986 | -105.4122
14.9023 | -105.4058
14.9061 | -105.4007
14.9099 | -105.3956
14.9137 | -105.3874
14.9174 | -105.3816
14.9212 | -105.3746
14.925 | -105.3648
14.9288 | -105.3533
14.9325 | -105.3431
14.9363 | -105.3335
14.9401 | -105.3227
14.9439 | -105.3089
14.9476 | -105.2944
14.9514 | -105.2821
14.9552 | -105.2695
14.959 | -105.2554
14.9627 | -105.2422
14.9665 | -105.2326
14.9703 | -105.2237
14.974 | -105.2149
14.9778 | -105.208
14.9816 | -105.2035
14.9854 | -105.1992
14.9891 | -105.1983
14.9929 | -105.2012
14.9967 | -105.2053
15.0005 | -105.2125
15.0042 | -105.2225
15.008 | -105.2352
15.0118 | -105.2502
15.0156 | -105.2687
15.0193 | -105.29
15.0231 | -105.3124
15.0269 | -105.3364
15.0307 | -105.3652
15.0344 | -105.3929
15.0382 | -105.4231
15.042 | -105.4522
15.0458 | -105.481
15.0495 | -105.5095
15.0533 | -105.5396
15.0571 | -105.569
15.0609 | -105.5974
15.0646 | -105.6275
15.0684 | -105.659
15.0722 | -105.6875
15.076 | -105.7146
15.0797 | -105.7395
15.0835 | -105.7614
15.0873 | -105.7842
15.0911 | -105.8033
15.0948 | -105.8181
15.0986 | -105.8314
15.1024 | -105.8444
15.1062 | -105.8553
15.1099 | -105.8686
15.1137 | -105.8756
15.1175 | -105.882
15.1213 | -105.8866
15.125 | -105.8901
15.1288 | -105.8895
15.1326 | -105.883
15.1364 | -105.8785
15.1401 | -105.8754
15.1439 | -105.8738
15.1477 | -105.8675
15.1515 | -105.8584
15.1552 | -105.8533
15.159 | -105.8478
15.1628 | -105.8422
15.1666 | -105.8366
15.1703 | -105.8295
15.1741 | -105.8221
15.1779 | -105.8161
15.1817 | -105.8097
15.1854 | -105.8014
15.1892 | -105.7916
15.193 | -105.784
15.1968 | -105.7799
15.2005 | -105.7739
15.2043 | -105.7647
15.2081 | -105.7557
15.2119 | -105.7477
15.2156 | -105.7343
15.2194 | -105.7197
15.2232 | -105.7084
15.227 | -105.7008
15.2307 | -105.6909
15.2345 | -105.682
15.2383 | -105.6735
15.2421 | -105.6645
15.2458 | -105.6495
15.2496 | -105.6391
15.2534 | -105.6286
15.2571 | -105.6161
15.2609 | -105.6041
15.2647 | -105.5933
15.2685 | -105.5811
15.2722 | -105.568
15.276 | -105.5557
15.2798 | -105.545
15.2836 | -105.5339
15.2873 | -105.522
15.2911 | -105.5129
15.2949 | -105.5062
15.2987 | -105.5006
15.3024 | -105.497
15.3062 | -105.4959
15.31 | -105.4953
15.3138 | -105.4952
15.3175 | -105.4965
15.3213 | -105.4998
15.3251 | -105.5024
15.3289 | -105.5055
15.3326 | -105.5064
15.3364 | -105.5111
15.3402 | -105.5165
15.344 | -105.5207
15.3477 | -105.5287
15.3515 | -105.5377
15.3553 | -105.5443
15.3591 | -105.5459
15.3628 | -105.5492
15.3666 | -105.5528
15.3704 | -105.5543
15.3742 | -105.554
15.3779 | -105.5527
15.3817 | -105.5468
15.3855 | -105.54
15.3893 | -105.5345
15.393 | -105.5291
15.3968 | -105.5236
15.4006 | -105.5161
15.4044 | -105.5084
15.4081 | -105.5018
15.4119 | -105.4895
15.4157 | -105.4711
15.4195 | -105.4577
15.4232 | -105.4488
15.427 | -105.4375
15.4308 | -105.4259
15.4346 | -105.4181
15.4383 | -105.4102
15.4421 | -105.3996
15.4459 | -105.3867
15.4497 | -105.3757
15.4534 | -105.3663
15.4572 | -105.3587
15.461 | -105.3531
15.4648 | -105.3495
15.4685 | -105.3449
15.4723 | -105.3415
15.4761 | -105.3404
15.4799 | -105.3406
15.4836 | -105.3397
15.4874 | -105.3387
15.4912 | -105.3401
15.495 | -105.3414
15.4987 | -105.3442
15.5025 | -105.3459
15.5063 | -105.3454
15.51 | -105.3466
15.5138 | -105.3515
15.5176 | -105.3571
15.5214 | -105.3626
15.5251 | -105.3652
15.5289 | -105.3668
15.5327 | -105.3713
15.5365 | -105.3761
15.5402 | -105.3825
15.544 | -105.3889
15.5478 | -105.3954
15.5516 | -105.3995
15.5553 | -105.4055
15.5591 | -105.4077
15.5629 | -105.4101
15.5667 | -105.4155
15.5704 | -105.4204
15.5742 | -105.4245
15.578 | -105.4289
15.5818 | -105.4368
15.5855 | -105.4417
15.5893 | -105.445
15.5931 | -105.4482
15.5969 | -105.4534
15.6006 | -105.4591
15.6044 | -105.4633
15.6082 | -105.4665
15.612 | -105.4697
15.6157 | -105.4703
15.6195 | -105.4662
15.6233 | -105.4648
15.6271 | -105.4656
15.6308 | -105.4666
15.6346 | -105.4698
15.6384 | -105.4741
15.6422 | -105.4718
15.6459 | -105.4675
15.6497 | -105.4638
15.6535 | -105.4603
15.6573 | -105.4567
15.661 | -105.4525
15.6648 | -105.4506
15.6686 | -105.4479
15.6724 | -105.4445
15.6761 | -105.4386
15.6799 | -105.4332
15.6837 | -105.4267
15.6875 | -105.4192
15.6912 | -105.4109
15.695 | -105.4001
15.6988 | -105.3872
15.7026 | -105.3752
15.7063 | -105.3663
15.7101 | -105.3567
15.7139 | -105.3477
15.7177 | -105.3385
15.7214 | -105.3287
15.7252 | -105.322
15.729 | -105.3153
15.7328 | -105.3068
15.7365 | -105.2954
15.7403 | -105.2866
15.7441 | -105.2791
15.7479 | -105.2753
15.7516 | -105.2719
15.7554 | -105.2642
15.7592 | -105.2571
15.763 | -105.249
15.7667 | -105.2403
15.7705 | -105.234
15.7743 | -105.2297
15.7781 | -105.2254
15.7818 | -105.2226
15.7856 | -105.2212
15.7894 | -105.2229
15.7931 | -105.2178
15.7969 | -105.2148
15.8007 | -105.2138
15.8045 | -105.2093
15.8082 | -105.2044
15.812 | -105.2008
15.8158 | -105.1995
15.8196 | -105.2006
15.8233 | -105.2001
15.8271 | -105.2013
15.8309 | -105.2042
15.8347 | -105.2067
15.8384 | -105.2099
15.8422 | -105.2121
15.846 | -105.2159
15.8498 | -105.2192
15.8535 | -105.2215
15.8573 | -105.2233
15.8611 | -105.2265
15.8649 | -105.2327
15.8686 | -105.2401
15.8724 | -105.2477
15.8762 | -105.2541
15.88 | -105.2589
15.8837 | -105.2627
15.8875 | -105.2667
15.8913 | -105.2703
15.8951 | -105.2758
15.8988 | -105.2827
15.9026 | -105.2927
15.9064 | -105.2992
15.9102 | -105.3068
15.9139 | -105.3169
15.9177 | -105.3291
15.9215 | -105.3405
15.9253 | -105.3532
15.929 | -105.3658
15.9328 | -105.3754
15.9366 | -105.3841
15.9404 | -105.3914
15.9441 | -105.3992
15.9479 | -105.4076
15.9517 | -105.4162
15.9555 | -105.4255
15.9592 | -105.4351
15.963 | -105.4433
15.9668 | -105.4508
15.9706 | -105.4583
15.9743 | -105.4649
15.9781 | -105.4727
15.9819 | -105.4841
15.9857 | -105.4908
15.9894 | -105.4935
15.9932 | -105.4951
15.997 | -105.4987
16.0008 | -105.5042
16.0045 | -105.5068
16.0083 | -105.5053
16.0121 | -105.5017
16.0159 | -105.4998
16.0196 | -105.4981
16.0234 | -105.4984
16.0272 | -105.4965
16.031 | -105.4957
16.0347 | -105.4979
16.0385 | -105.4991
16.0423 | -105.4982
16.0461 | -105.4959
16.0498 | -105.4916
16.0536 | -105.4894
16.0574 | -105.4865
16.0611 | -105.4821
16.0649 | -105.4791
16.0687 | -105.4731
16.0725 | -105.4688
16.0762 | -105.4653
16.08 | -105.4625
16.0838 | -105.4615
16.0876 | -105.4605
16.0913 | -105.4574
16.0951 | -105.4524
16.0989 | -105.4489
16.1027 | -105.444
16.1064 | -105.4416
16.1102 | -105.437
16.114 | -105.4323
16.1178 | -105.4273
16.1215 | -105.4215
16.1253 | -105.4142
16.1291 | -105.4065
16.1329 | -105.3992
16.1366 | -105.3958
16.1404 | -105.3897
16.1442 | -105.3808
16.148 | -105.3708
16.1517 | -105.363
16.1555 | -105.3527
16.1593 | -105.3426
16.1631 | -105.3358
16.1668 | -105.3302
16.1706 | -105.3242
16.1744 | -105.3189
16.1782 | -105.3119
16.1819 | -105.3039
16.1857 | -105.2978
16.1895 | -105.2943
16.1933 | -105.2923
16.197 | -105.29
16.2008 | -105.2877
16.2046 | -105.2861
16.2084 | -105.2826
16.2121 | -105.2789
16.2159 | -105.2815
16.2197 | -105.2866
16.2235 | -105.292
16.2272 | -105.2946
16.231 | -105.2968
16.2348 | -105.3019
16.2386 | -105.3101
16.2423 | -105.3184
16.2461 | -105.3249
16.2499 | -105.3299
16.2537 | -105.3377
16.2574 | -105.3446
16.2612 | -105.3525
16.265 | -105.3616
16.2688 | -105.3706
16.2725 | -105.3807
16.2763 | -105.3933
16.2801 | -105.406
16.2839 | -105.4138
16.2876 | -105.422
16.2914 | -105.4342
16.2952 | -105.4442
16.299 | -105.4548
16.3027 | -105.4647
16.3065 | -105.4754
16.3103 | -105.4837
16.3141 | -105.4897
16.3178 | -105.4956
16.3216 | -105.506
16.3254 | -105.5172
16.3291 | -105.5244
16.3329 | -105.5327
16.3367 | -105.5393
16.3405 | -105.5447
16.3442 | -105.5524
16.348 | -105.5608
16.3518 | -105.5675
16.3556 | -105.5724
16.3593 | -105.5782
16.3631 | -105.5818
16.3669 | -105.5823
16.3707 | -105.581
16.3744 | -105.5799
16.3782 | -105.5781
16.382 | -105.5773
16.3858 | -105.5758
16.3895 | -105.5752
16.3933 | -105.5723
16.3971 | -105.5669
16.4009 | -105.5618
16.4046 | -105.5553
16.4084 | -105.5494
16.4122 | -105.5449
16.416 | -105.5387
16.4197 | -105.5318
16.4235 | -105.5229
16.4273 | -105.5133
16.4311 | -105.5055
16.4348 | -105.499
16.4386 | -105.4901
16.4424 | -105.4826
16.4462 | -105.4761
16.4499 | -105.4659
16.4537 | -105.4565
16.4575 | -105.4492
16.4613 | -105.4414
16.465 | -105.4339
16.4688 | -105.4277
16.4726 | -105.42
16.4764 | -105.4123
16.4801 | -105.4058
16.4839 | -105.3998
16.4877 | -105.3929
16.4915 | -105.3905
16.4952 | -105.3849
16.499 | -105.3784
16.5028 | -105.3738
16.5066 | -105.3721
16.5103 | -105.3703
16.5141 | -105.3688
16.5179 | -105.3688
16.5217 | -105.3683
16.5254 | -105.3661
16.5292 | -105.3665
16.533 | -105.3678
16.5368 | -105.3708
16.5405 | -105.3728
16.5443 | -105.3752
16.5481 | -105.3762
16.5519 | -105.3778
16.5556 | -105.38
16.5594 | -105.3836
16.5632 | -105.3883
16.567 | -105.394
16.5707 | -105.3997
16.5745 | -105.4086
16.5783 | -105.4181
16.5821 | -105.4248
16.5858 | -105.4316
16.5896 | -105.4417
16.5934 | -105.4533
16.5971 | -105.4643
16.6009 | -105.4779
16.6047 | -105.4922
16.6085 | -105.5081
16.6122 | -105.5248
16.616 | -105.5386
16.6198 | -105.5494
16.6236 | -105.5609
16.6273 | -105.5736
16.6311 | -105.5873
16.6349 | -105.6003
16.6387 | -105.6105
16.6424 | -105.6222
16.6462 | -105.6328
16.65 | -105.6432
16.6538 | -105.6551
16.6575 | -105.664
16.6613 | -105.6724
16.6651 | -105.6832
16.6689 | -105.6933
16.6726 | -105.701
16.6764 | -105.7118
16.6802 | -105.7224
16.684 | -105.7303
16.6877 | -105.7397
16.6915 | -105.7492
16.6953 | -105.755
16.6991 | -105.7624
16.7028 | -105.7692
16.7066 | -105.7765
16.7104 | -105.7827
16.7142 | -105.7894
16.7179 | -105.7976
16.7217 | -105.8063
16.7255 | -105.8134
16.7293 | -105.8149
16.733 | -105.8196
16.7368 | -105.8267
16.7406 | -105.8285
16.7444 | -105.8248
16.7481 | -105.8243
16.7519 | -105.8269
16.7557 | -105.8296
16.7595 | -105.8313
16.7632 | -105.8307
16.767 | -105.8283
16.7708 | -105.824
16.7746 | -105.8185
16.7783 | -105.8121
16.7821 | -105.8074
16.7859 | -105.8062
16.7897 | -105.8019
16.7934 | -105.7924
16.7972 | -105.7843
16.801 | -105.7774
16.8048 | -105.7677
16.8085 | -105.7572
16.8123 | -105.7458
16.8161 | -105.7339
16.8199 | -105.7214
16.8236 | -105.7075
16.8274 | -105.6947
16.8312 | -105.6817
16.835 | -105.6667
16.8387 | -105.6531
16.8425 | -105.6385
16.8463 | -105.6249
16.8501 | -105.6128
16.8538 | -105.6045
16.8576 | -105.5971
16.8614 | -105.5873
16.8652 | -105.5774
16.8689 | -105.5697
16.8727 | -105.5627
16.8765 | -105.5568
16.8802 | -105.5564
16.884 | -105.5547
16.8878 | -105.5497
16.8916 | -105.5441
16.8953 | -105.5423
16.8991 | -105.5444
16.9029 | -105.5457
16.9067 | -105.5455
16.9104 | -105.5454
16.9142 | -105.5481
16.918 | -105.5491
16.9218 | -105.5528
16.9255 | -105.558
16.9293 | -105.5624
16.9331 | -105.5651
16.9369 | -105.5711
16.9406 | -105.5776
16.9444 | -105.5821
16.9482 | -105.5856
16.952 | -105.5916
16.9557 | -105.5973
16.9595 | -105.5998
16.9633 | -105.601
16.9671 | -105.6026
16.9708 | -105.6059
16.9746 | -105.6083
16.9784 | -105.6085
16.9822 | -105.6102
16.9859 | -105.6113
16.9897 | -105.6106
16.9935 | -105.6096
16.9973 | -105.61
17.001 | -105.6104
17.0048 | -105.6098
17.0086 | -105.6082
17.0124 | -105.6074
17.0161 | -105.604
17.0199 | -105.6022
17.0237 | -105.6012
17.0275 | -105.6007
17.0312 | -105.6016
17.035 | -105.6028
17.0388 | -105.6037
17.0426 | -105.6054
17.0463 | -105.6066
17.0501 | -105.6075
17.0539 | -105.6094
17.0577 | -105.6127
17.0614 | -105.6162
17.0652 | -105.6195
17.069 | -105.6227
17.0728 | -105.6265
17.0765 | -105.6319
17.0803 | -105.6364
17.0841 | -105.6417
17.0879 | -105.6446
17.0916 | -105.6465
17.0954 | -105.6483
17.0992 | -105.6518
17.103 | -105.6584
17.1067 | -105.6638
17.1105 | -105.6684
17.1143 | -105.6747
17.1181 | -105.6826
17.1218 | -105.6888
17.1256 | -105.6948
17.1294 | -105.6984
17.1332 | -105.7055
17.1369 | -105.712
17.1407 | -105.7166
17.1445 | -105.7208
17.1483 | -105.7265
17.152 | -105.7314
17.1558 | -105.7363
17.1596 | -105.7418
17.1633 | -105.7459
17.1671 | -105.7479
17.1709 | -105.7502
17.1747 | -105.7525
17.1784 | -105.7543
17.1822 | -105.7547
17.186 | -105.7547
17.1898 | -105.7578
17.1935 | -105.7587
17.1973 | -105.7562
17.2011 | -105.7523
17.2049 | -105.7499
17.2086 | -105.7467
17.2124 | -105.7426
17.2162 | -105.7379
17.22 | -105.733
17.2237 | -105.7287
17.2275 | -105.7253
17.2313 | -105.7238
17.2351 | -105.718
17.2388 | -105.7105
17.2426 | -105.704
17.2464 | -105.6992
17.2502 | -105.6984
17.2539 | -105.6966
17.2577 | -105.6931
17.2615 | -105.6921
17.2653 | -105.6886
17.269 | -105.6861
17.2728 | -105.6842
17.2766 | -105.6847
17.2804 | -105.6859
17.2841 | -105.686
17.2879 | -105.6867
17.2917 | -105.6897
17.2955 | -105.6936
17.2992 | -105.6976
17.303 | -105.7022
17.3068 | -105.7089
17.3106 | -105.7139
17.3143 | -105.7159
17.3181 | -105.7193
17.3219 | -105.7243
17.3257 | -105.7304
17.3294 | -105.7397
17.3332 | -105.7508
17.337 | -105.7592
17.3408 | -105.7682
17.3445 | -105.7768
17.3483 | -105.7842
17.3521 | -105.7918
17.3559 | -105.7992
17.3596 | -105.806
17.3634 | -105.8121
17.3672 | -105.8185
17.371 | -105.8225
17.3747 | -105.8252
17.3785 | -105.8282
17.3823 | -105.8317
17.3861 | -105.8338
17.3898 | -105.8355
17.3936 | -105.8378
17.3974 | -105.8413
17.4012 | -105.8446
17.4049 | -105.8458
17.4087 | -105.8482
17.4125 | -105.8509
17.4163 | -105.8534
17.42 | -105.8578
17.4238 | -105.8645
17.4276 | -105.8706
17.4313 | -105.8771
17.4351 | -105.8824
17.4389 | -105.8877
17.4427 | -105.8947
17.4464 | -105.9039
17.4502 | -105.912
17.454 | -105.917
17.4578 | -105.9235
17.4615 | -105.9328
17.4653 | -105.9417
17.4691 | -105.9501
17.4729 | -105.9582
17.4766 | -105.9666
17.4804 | -105.9752
17.4842 | -105.9856
17.488 | -105.9935
17.4917 | -105.9983
17.4955 | -106.0021
17.4993 | -106.0024
17.5031 | -106.0043
17.5068 | -106.0077
17.5106 | -106.0076
17.5144 | -106.0061
17.5182 | -106.0007
17.5219 | -105.9943
17.5257 | -105.9863
17.5295 | -105.9793
17.5333 | -105.9719
17.537 | -105.9634
17.5408 | -105.9522
17.5446 | -105.938
17.5484 | -105.9223
17.5521 | -105.9087
17.5559 | -105.8947
17.5597 | -105.8793
17.5635 | -105.8648
17.5672 | -105.8498
17.571 | -105.8299
17.5748 | -105.8099
17.5786 | -105.7909
17.5823 | -105.7721
17.5861 | -105.7573
17.5899 | -105.7395
17.5937 | -105.719
17.5974 | -105.7026
17.6012 | -105.6868
17.605 | -105.6726
17.6088 | -105.6608
17.6125 | -105.6487
17.6163 | -105.637
17.6201 | -105.626
17.6239 | -105.6175
17.6276 | -105.6087
17.6314 | -105.5975
17.6352 | -105.584
17.639 | -105.5737
17.6427 | -105.5641
17.6465 | -105.5532
17.6503 | -105.5416
17.6541 | -105.532
17.6578 | -105.5238
17.6616 | -105.5142
17.6654 | -105.5048
17.6692 | -105.4961
17.6729 | -105.4871
17.6767 | -105.4786
17.6805 | -105.4709
17.6842 | -105.4624
17.688 | -105.4561
17.6918 | -105.4512
17.6956 | -105.4484
17.6993 | -105.4462
17.7031 | -105.4445
17.7069 | -105.4445
17.7107 | -105.4446
17.7144 | -105.4433
17.7182 | -105.4422
17.722 | -105.4424
17.7258 | -105.4462
17.7295 | -105.4505
17.7333 | -105.4531
17.7371 | -105.4542
17.7409 | -105.4594
17.7446 | -105.4665
17.7484 | -105.4753
17.7522 | -105.483
17.756 | -105.491
17.7597 | -105.4996
17.7635 | -105.507
17.7673 | -105.5149
17.7711 | -105.5226
17.7748 | -105.5299
17.7786 | -105.5342
17.7824 | -105.5383
17.7862 | -105.5419
17.7899 | -105.5463
17.7937 | -105.5525
17.7975 | -105.5583
17.8013 | -105.5633
17.805 | -105.5666
17.8088 | -105.5703
17.8126 | -105.5731
17.8164 | -105.5756
17.8201 | -105.5778
17.8239 | -105.5829
17.8277 | -105.5857
17.8315 | -105.5844
17.8352 | -105.5837
17.839 | -105.5814
17.8428 | -105.5804
17.8466 | -105.5814
17.8503 | -105.5791
17.8541 | -105.5772
17.8579 | -105.576
17.8617 | -105.5719
17.8654 | -105.5657
17.8692 | -105.5595
17.873 | -105.5549
17.8768 | -105.5492
17.8805 | -105.5426
17.8843 | -105.5387
17.8881 | -105.5357
17.8919 | -105.5293
17.8956 | -105.5227
17.8994 | -105.5162
17.9032 | -105.5091
17.907 | -105.4999
17.9107 | -105.4912
17.9145 | -105.484
17.9183 | -105.4754
17.9221 | -105.4656
17.9258 | -105.4567
17.9296 | -105.4485
17.9334 | -105.4409
17.9372 | -105.434
17.9409 | -105.4246
17.9447 | -105.4113
17.9485 | -105.4002
17.9523 | -105.391
17.956 | -105.3832
17.9598 | -105.3762
17.9636 | -105.3681
17.9673 | -105.3632
17.9711 | -105.3568
17.9749 | -105.3539
17.9787 | -105.3488
17.9824 | -105.3415
17.9862 | -105.3372
17.99 | -105.3364
17.9938 | -105.3363
17.9975 | -105.3345
18.0013 | -105.3335
18.0051 | -105.3333
18.0089 | -105.3339
18.0126 | -105.3349
18.0164 | -105.3375
18.0202 | -105.3427
18.024 | -105.3486
18.0277 | -105.3547
18.0315 | -105.3613
18.0353 | -105.3719
18.0391 | -105.3823
18.0428 | -105.3912
18.0466 | -105.4003
18.0504 | -105.4083
18.0542 | -105.4174
18.0579 | -105.4282
18.0617 | -105.4393
18.0655 | -105.4491
18.0693 | -105.4611
18.073 | -105.4748
18.0768 | -105.4859
18.0806 | -105.4963
18.0844 | -105.5086
18.0881 | -105.5221
18.0919 | -105.535
18.0957 | -105.5484
18.0995 | -105.5622
18.1032 | -105.5735
18.107 | -105.5835
18.1108 | -105.5935
18.1146 | -105.6027
18.1183 | -105.6125
18.1221 | -105.6267
18.1259 | -105.6427
18.1297 | -105.6544
18.1334 | -105.664
18.1372 | -105.6737
18.141 | -105.6843
18.1448 | -105.6927
18.1485 | -105.7017
18.1523 | -105.7096
18.1561 | -105.7174
18.1599 | -105.7255
18.1636 | -105.7368
18.1674 | -105.744
18.1712 | -105.7486
18.175 | -105.7546
18.1787 | -105.7606
18.1825 | -105.7661
18.1863 | -105.7695
18.1901 | -105.7695
18.1938 | -105.7689
18.1976 | -105.7721
18.2014 | -105.7755
18.2052 | -105.7757
18.2089 | -105.7753
18.2127 | -105.7718
18.2165 | -105.7696
18.2203 | -105.7674
18.224 | -105.7644
18.2278 | -105.7613
18.2316 | -105.7549
18.2354 | -105.7446
18.2391 | -105.7371
18.2429 | -105.7303
18.2467 | -105.7206
18.2504 | -105.7075
18.2542 | -105.6961
18.258 | -105.6858
18.2618 | -105.675
18.2655 | -105.6625
18.2693 | -105.6509
18.2731 | -105.6382
18.2769 | -105.6229
18.2806 | -105.6061
18.2844 | -105.5901
18.2882 | -105.5761
18.292 | -105.5627
18.2957 | -105.546
18.2995 | -105.5288
18.3033 | -105.5127
18.3071 | -105.4944
18.3108 | -105.478
18.3146 | -105.4619
18.3184 | -105.4446
18.3222 | -105.4272
18.3259 | -105.4117
18.3297 | -105.3945
18.3335 | -105.3755
18.3373 | -105.3583
18.341 | -105.3413
18.3448 | -105.321
18.3486 | -105.2989
18.3524 | -105.2804
18.3561 | -105.2617
18.3599 | -105.2416
18.3637 | -105.2238
18.3675 | -105.2062
18.3712 | -105.1899
18.375 | -105.1723
18.3788 | -105.154
18.3826 | -105.135
18.3863 | -105.1152
18.3901 | -105.0978
18.3939 | -105.0815
18.3977 | -105.0668
18.4014 | -105.0508
18.4052 | -105.0352
18.409 | -105.0189
18.4128 | -105.0033
18.4165 | -104.9881
18.4203 | -104.9726
18.4241 | -104.9549
18.4279 | -104.9392
18.4316 | -104.9266
18.4354 | -104.9132
18.4392 | -104.8982
18.443 | -104.8847
18.4467 | -104.8726
18.4505 | -104.861
18.4543 | -104.8486
18.4581 | -104.8333
18.4618 | -104.8198
18.4656 | -104.8075
18.4694 | -104.795
18.4732 | -104.7838
18.4769 | -104.7715
18.4807 | -104.7587
18.4845 | -104.7464
18.4883 | -104.7359
18.492 | -104.7267
18.4958 | -104.7207
18.4996 | -104.7149
18.5034 | -104.7096
18.5071 | -104.7058
18.5109 | -104.7041
18.5147 | -104.7067
18.5184 | -104.7076
18.5222 | -104.7078
18.526 | -104.7078
18.5298 | -104.7092
18.5335 | -104.7133
18.5373 | -104.72
18.5411 | -104.7287
18.5449 | -104.7383
18.5486 | -104.7453
18.5524 | -104.7515
18.5562 | -104.7598
18.56 | -104.7687
18.5637 | -104.7787
18.5675 | -104.7901
18.5713 | -104.8015
18.5751 | -104.8106
18.5788 | -104.818
18.5826 | -104.8245
18.5864 | -104.8307
18.5902 | -104.8347
18.5939 | -104.8413
18.5977 | -104.8465
18.6015 | -104.8502
18.6053 | -104.8532
18.609 | -104.8554
18.6128 | -104.8573
18.6166 | -104.8537
18.6204 | -104.8471
18.6241 | -104.8393
18.6279 | -104.8313
18.6317 | -104.8221
18.6355 | -104.8113
18.6392 | -104.8012
18.643 | -104.7929
18.6468 | -104.7838
18.6506 | -104.7725
18.6543 | -104.7607
18.6581 | -104.7489
18.6619 | -104.735
18.6657 | -104.7234
18.6694 | -104.7143
18.6732 | -104.7059
18.677 | -104.6978
18.6808 | -104.6902
18.6845 | -104.6832
18.6883 | -104.6782
18.6921 | -104.674
18.6959 | -104.6687
18.6996 | -104.664
18.7034 | -104.6623
18.7072 | -104.6584
18.711 | -104.6574
18.7147 | -104.6585
18.7185 | -104.6615
18.7223 | -104.6675
18.7261 | -104.6736
18.7298 | -104.6795
18.7336 | -104.6855
18.7374 | -104.6939
18.7412 | -104.7027
18.7449 | -104.7115
18.7487 | -104.7194
18.7525 | -104.7286
18.7563 | -104.7384
18.76 | -104.7459
18.7638 | -104.7527
18.7676 | -104.7599
18.7714 | -104.7681
18.7751 | -104.7763
18.7789 | -104.7854
18.7827 | -104.7931
18.7864 | -104.7985
18.7902 | -104.8039
18.794 | -104.8089
18.7978 | -104.8145
18.8015 | -104.819
18.8053 | -104.8216
18.8091 | -104.825
18.8129 | -104.8293
18.8166 | -104.8351
18.8204 | -104.8416
18.8242 | -104.8472
18.828 | -104.8532
18.8317 | -104.8587
18.8355 | -104.8633
18.8393 | -104.8701
18.8431 | -104.8784
18.8468 | -104.886
18.8506 | -104.8942
18.8544 | -104.9011
18.8582 | -104.909
18.8619 | -104.9145
18.8657 | -104.9222
18.8695 | -104.9325
18.8733 | -104.9447
18.877 | -104.9568
18.8808 | -104.9693
18.8846 | -104.9828
18.8884 | -104.997
18.8921 | -105.0113
18.8959 | -105.0228
18.8997 | -105.0357
18.9035 | -105.0478
18.9072 | -105.0593
18.911 | -105.0685
18.9148 | -105.0774
18.9186 | -105.0877
18.9223 | -105.0966
18.9261 | -105.1032
18.9299 | -105.1117
18.9337 | -105.1178
18.9374 | -105.1212
18.9412 | -105.1228
18.945 | -105.1246
18.9488 | -105.1255
18.9525 | -105.1259
18.9563 | -105.1246
18.9601 | -105.122
18.9639 | -105.1189
18.9676 | -105.1155
18.9714 | -105.1102
18.9752 | -105.1077
18.979 | -105.0996
18.9827 | -105.0882
18.9865 | -105.0798
18.9903 | -105.0716
18.9941 | -105.0634
18.9978 | -105.0529
19.0016 | -105.042
19.0054 | -105.0313
19.0092 | -105.0233
19.0129 | -105.0158
19.0167 | -105.0079
19.0205 | -104.9995
19.0243 | -104.9924
19.028 | -104.9855
19.0318 | -104.978
19.0356 | -104.9721
19.0394 | -104.9653
19.0431 | -104.9593
19.0469 | -104.9562
19.0507 | -104.9535
19.0544 | -104.9504
19.0582 | -104.9554
19.062 | -104.9561
19.0658 | -104.9561
19.0695 | -104.9584
19.0733 | -104.9623
19.0771 | -104.965
19.0809 | -104.9668
19.0846 | -104.9697
19.0884 | -104.972
19.0922 | -104.9747
19.096 | -104.978
19.0997 | -104.9836
19.1035 | -104.9892
19.1073 | -104.9954
19.1111 | -105.0009
19.1148 | -105.0065
19.1186 | -105.014
19.1224 | -105.0198
19.1262 | -105.0269
19.1299 | -105.036
19.1337 | -105.0433
19.1375 | -105.0477
19.1413 | -105.0512
19.145 | -105.0555
19.1488 | -105.0614
19.1526 | -105.0666
19.1564 | -105.0725
19.1601 | -105.0795
19.1639 | -105.086
19.1677 | -105.0907
19.1715 | -105.0965
19.1752 | -105.1022
19.179 | -105.1079
19.1828 | -105.1145
19.1866 | -105.1194
19.1903 | -105.1249
19.1941 | -105.1316
19.1979 | -105.1407
19.2017 | -105.1474
19.2054 | -105.1534
19.2092 | -105.1587
19.213 | -105.1653
19.2168 | -105.1717
19.2205 | -105.1771
19.2243 | -105.1823
19.2281 | -105.1871
19.2319 | -105.194
19.2356 | -105.2011
19.2394 | -105.2065
19.2432 | -105.2099
19.247 | -105.2119
19.2507 | -105.2149
19.2545 | -105.2196
19.2583 | -105.2174
19.2621 | -105.2146
19.2658 | -105.2135
19.2696 | -105.2111
19.2734 | -105.2119
19.2772 | -105.2097
19.2809 | -105.2049
19.2847 | -105.1993
19.2885 | -105.1929
19.2923 | -105.1856
19.296 | -105.1774
19.2998 | -105.1698
19.3036 | -105.1635
19.3074 | -105.1546
19.3111 | -105.1438
19.3149 | -105.1333
19.3187 | -105.1196
19.3224 | -105.104
19.3262 | -105.0913
19.33 | -105.0748
19.3338 | -105.0561
19.3375 | -105.0396
19.3413 | -105.0218
19.3451 | -105.003
19.3489 | -104.9856
19.3526 | -104.9695
19.3564 | -104.9556
19.3602 | -104.9423
19.364 | -104.9299
19.3677 | -104.9169
19.3715 | -104.9017
19.3753 | -104.8854
19.3791 | -104.8706
19.3828 | -104.8567
19.3866 | -104.8444
19.3904 | -104.8336
19.3942 | -104.8249
19.3979 | -104.8164
19.4017 | -104.808
19.4055 | -104.8012
19.4093 | -104.7972
19.413 | -104.7952
19.4168 | -104.7942
19.4206 | -104.7913
19.4244 | -104.7888
19.4281 | -104.7863
19.4319 | -104.7856
19.4357 | -104.788
19.4395 | -104.7908
19.4432 | -104.7955
19.447 | -104.7993
19.4508 | -104.8061
19.4546 | -104.8123
19.4583 | -104.818
19.4621 | -104.8288
19.4659 | -104.8384
19.4697 | -104.8488
19.4734 | -104.8593
19.4772 | -104.8709
19.481 | -104.8823
19.4848 | -104.8944
19.4885 | -104.9078
19.4923 | -104.9213
19.4961 | -104.9348
19.4999 | -104.9476
19.5036 | -104.9606
19.5074 | -104.97
19.5112 | -104.9794
19.515 | -104.9895
19.5187 | -105.0011
19.5225 | -105.0105
19.5263 | -105.0197
19.5301 | -105.0283
19.5338 | -105.0369
19.5376 | -105.0411
19.5414 | -105.0411
19.5452 | -105.0433
19.5489 | -105.042
19.5527 | -105.042
19.5565 | -105.0424
19.5603 | -105.0415
19.564 | -105.0389
19.5678 | -105.0325
19.5716 | -105.0251
19.5754 | -105.0174
19.5791 | -105.007
19.5829 | -104.9973
19.5867 | -104.9884
19.5905 | -104.9787
19.5942 | -104.9706
19.598 | -104.9626
19.6018 | -104.9508
19.6055 | -104.94
19.6093 | -104.9295
19.6131 | -104.9197
19.6169 | -104.9083
19.6206 | -104.9006
19.6244 | -104.8913
19.6282 | -104.8794
19.632 | -104.8705
19.6357 | -104.8608
19.6395 | -104.8505
19.6433 | -104.841
19.6471 | -104.8339
19.6508 | -104.8291
19.6546 | -104.8249
19.6584 | -104.8221
19.6622 | -104.8188
19.6659 | -104.8161
19.6697 | -104.8138
19.6735 | -104.8094
19.6773 | -104.8094
19.681 | -104.8119
19.6848 | -104.8127
19.6886 | -104.8146
19.6924 | -104.8177
19.6961 | -104.8218
19.6999 | -104.8241
19.7037 | -104.8292
19.7075 | -104.8383
19.7112 | -104.8456
19.715 | -104.8505
19.7188 | -104.8565
19.7226 | -104.8633
19.7263 | -104.8708
19.7301 | -104.8772
19.7339 | -104.8851
19.7377 | -104.8931
19.7414 | -104.8975
19.7452 | -104.9024
19.749 | -104.9082
19.7528 | -104.9127
19.7565 | -104.9183
19.7603 | -104.9243
19.7641 | -104.93
19.7679 | -104.9328
19.7716 | -104.9329
19.7754 | -104.9332
19.7792 | -104.9357
19.783 | -104.9364
19.7867 | -104.9366
19.7905 | -104.9367
19.7943 | -104.9373
19.7981 | -104.9368
19.8018 | -104.9372
19.8056 | -104.9369
19.8094 | -104.936
19.8132 | -104.9366
19.8169 | -104.9382
19.8207 | -104.9389
19.8245 | -104.9372
19.8283 | -104.9334
19.832 | -104.9311
19.8358 | -104.9281
19.8396 | -104.9253
19.8434 | -104.9243
19.8471 | -104.9262
19.8509 | -104.9261
19.8547 | -104.9247
19.8585 | -104.9254
19.8622 | -104.925
19.866 | -104.9238
19.8698 | -104.9212
19.8735 | -104.9182
19.8773 | -104.917
19.8811 | -104.9143
19.8849 | -104.913
19.8886 | -104.911
19.8924 | -104.9116
19.8962 | -104.9132
19.9 | -104.9133
19.9037 | -104.9121
19.9075 | -104.9099
19.9113 | -104.9071
19.9151 | -104.9049
19.9188 | -104.9033
19.9226 | -104.9021
19.9264 | -104.9023
19.9302 | -104.8989
19.9339 | -104.8962
19.9377 | -104.8947
19.9415 | -104.8906
19.9453 | -104.8854
19.949 | -104.8806
19.9528 | -104.8775
19.9566 | -104.876
19.9604 | -104.8723
19.9641 | -104.8659
19.9679 | -104.8588
19.9717 | -104.8526
19.9755 | -104.8454
19.9792 | -104.838
19.983 | -104.8327
19.9868 | -104.8233
19.9906 | -104.8135
19.9943 | -104.8048
19.9981 | -104.7964
20.0019 | -104.7871
20.0057 | -104.7782
20.0094 | -104.7669
20.0132 | -104.7536
20.017 | -104.7403
20.0208 | -104.7268
20.0245 | -104.7143
20.0283 | -104.7024
20.0321 | -104.6903
20.0359 | -104.6787
20.0396 | -104.6665
20.0434 | -104.6555
20.0472 | -104.6444
20.051 | -104.6349
20.0547 | -104.6264
20.0585 | -104.6173
20.0623 | -104.6064
20.0661 | -104.5967
20.0698 | -104.5888
20.0736 | -104.5812
20.0774 | -104.5734
20.0812 | -104.5658
20.0849 | -104.5595
20.0887 | -104.554
20.0925 | -104.5478
20.0963 | -104.5419
20.1 | -104.5369
20.1038 | -104.5341
20.1076 | -104.5324
20.1114 | -104.5304
20.1151 | -104.5267
20.1189 | -104.5241
20.1227 | -104.5247
20.1265 | -104.5238
20.1302 | -104.5231
20.134 | -104.5233
20.1378 | -104.5261
20.1415 | -104.5273
20.1453 | -104.5321
20.1491 | -104.5337
20.1529 | -104.5365
20.1566 | -104.5401
20.1604 | -104.5445
20.1642 | -104.5484
20.168 | -104.5534
20.1717 | -104.5584
20.1755 | -104.5646
20.1793 | -104.5705
20.1831 | -104.5753
20.1868 | -104.581
20.1906 | -104.5886
20.1944 | -104.5977
20.1982 | -104.6081
20.2019 | -104.6164
20.2057 | -104.6252
20.2095 | -104.6349
20.2133 | -104.6445
20.217 | -104.656
20.2208 | -104.6672
20.2246 | -104.6766
20.2284 | -104.6868
20.2321 | -104.6969
20.2359 | -104.7056
20.2397 | -104.7108
20.2435 | -104.7162
20.2472 | -104.7192
20.251 | -104.7249
20.2548 | -104.7323
20.2586 | -104.7383
20.2623 | -104.7435
20.2661 | -104.7483
20.2699 | -104.7528
20.2737 | -104.7577
20.2774 | -104.7614
20.2812 | -104.7635
20.285 | -104.7654
20.2888 | -104.7676
20.2925 | -104.7694
20.2963 | -104.7706
20.3001 | -104.7743
20.3039 | -104.7784
20.3076 | -104.7808
20.3114 | -104.7818
20.3152 | -104.7809
20.319 | -104.7784
20.3227 | -104.7753
20.3265 | -104.7722
20.3303 | -104.7692
20.3341 | -104.7668
20.3378 | -104.766
20.3416 | -104.7647
20.3454 | -104.7635
20.3492 | -104.7628
20.3529 | -104.7619
20.3567 | -104.7606
20.3605 | -104.7567
20.3643 | -104.7535
20.368 | -104.7497
20.3718 | -104.7456
20.3756 | -104.7432
20.3794 | -104.7399
20.3831 | -104.7372
20.3869 | -104.7338
20.3907 | -104.7307
20.3945 | -104.7277
20.3982 | -104.7267
20.402 | -104.7274
20.4058 | -104.7254
20.4095 | -104.7215
20.4133 | -104.7199
20.4171 | -104.719
20.4209 | -104.7187
20.4246 | -104.7158
20.4284 | -104.7142
20.4322 | -104.7148
20.436 | -104.7164
20.4397 | -104.7166
20.4435 | -104.7168
20.4473 | -104.7172
20.4511 | -104.7179
20.4548 | -104.7174
20.4586 | -104.7176
20.4624 | -104.7181
20.4662 | -104.7189
20.4699 | -104.7176
20.4737 | -104.7141
20.4775 | -104.7097
20.4813 | -104.7067
20.485 | -104.7037
20.4888 | -104.7003
20.4926 | -104.6962
20.4964 | -104.691
20.5001 | -104.6867
20.5039 | -104.6814
20.5077 | -104.6743
20.5115 | -104.6672
20.5152 | -104.661
20.519 | -104.6539
20.5228 | -104.65
20.5266 | -104.6463
20.5303 | -104.6429
20.5341 | -104.6366
20.5379 | -104.6307
20.5417 | -104.6215
20.5454 | -104.6121
20.5492 | -104.6039
20.553 | -104.5946
20.5568 | -104.5844
20.5605 | -104.5765
20.5643 | -104.5689
20.5681 | -104.5583
20.5719 | -104.549
20.5756 | -104.5411
20.5794 | -104.5335
20.5832 | -104.5239
20.587 | -104.5168
20.5907 | -104.5104
20.5945 | -104.5047
20.5983 | -104.5017
20.6021 | -104.4982
20.6058 | -104.4968
20.6096 | -104.4945
20.6134 | -104.4923
20.6172 | -104.4896
20.6209 | -104.4893
20.6247 | -104.4889
20.6285 | -104.4886
20.6323 | -104.4899
20.636 | -104.492
20.6398 | -104.494
20.6436 | -104.4964
20.6474 | -104.5014
20.6511 | -104.5081
20.6549 | -104.5133
20.6587 | -104.5205
20.6625 | -104.526
20.6662 | -104.5314
20.67 | -104.54
20.6738 | -104.5483
20.6776 | -104.5564
20.6813 | -104.5632
20.6851 | -104.5708
20.6889 | -104.5791
20.6926 | -104.5866
20.6964 | -104.5915
20.7002 | -104.5993
20.704 | -104.6068
20.7077 | -104.6114
20.7115 | -104.6143
20.7153 | -104.619
20.7191 | -104.625
20.7228 | -104.6316
20.7266 | -104.637
20.7304 | -104.6413
20.7342 | -104.6455
20.7379 | -104.6497
20.7417 | -104.6522
20.7455 | -104.6548
20.7493 | -104.6585
20.753 | -104.6622
20.7568 | -104.6651
20.7606 | -104.6679
20.7644 | -104.6717
20.7681 | -104.6789
20.7719 | -104.6867
20.7757 | -104.6914
20.7795 | -104.694
20.7832 | -104.6953
20.787 | -104.6987
20.7908 | -104.702
20.7946 | -104.7057
20.7983 | -104.7088
20.8021 | -104.7122
20.8059 | -104.7151
20.8097 | -104.7169
20.8134 | -104.7187
20.8172 | -104.7176
20.821 | -104.7163
20.8248 | -104.7169
20.8285 | -104.7196
20.8323 | -104.7216
20.8361 | -104.7211
20.8399 | -104.718
20.8436 | -104.7153
20.8474 | -104.7125
20.8512 | -104.71
20.855 | -104.7055
20.8587 | -104.7002
20.8625 | -104.6964
20.8663 | -104.6935
20.8701 | -104.6899
20.8738 | -104.6853
20.8776 | -104.6807
20.8814 | -104.6749
20.8852 | -104.668
20.8889 | -104.6592
20.8927 | -104.6503
20.8965 | -104.64
20.9003 | -104.6304
20.904 | -104.6231
20.9078 | -104.6177
20.9116 | -104.6122
20.9154 | -104.604
20.9191 | -104.5961
20.9229 | -104.5875
20.9267 | -104.5799
20.9305 | -104.5719
20.9342 | -104.564
20.938 | -104.5562
20.9418 | -104.5465
20.9456 | -104.5362
20.9493 | -104.5286
20.9531 | -104.522
20.9569 | -104.5139
20.9606 | -104.506
20.9644 | -104.4976
20.9682 | -104.4898
20.972 | -104.4813
20.9757 | -104.4726
20.9795 | -104.4644
20.9833 | -104.4555
20.9871 | -104.4459
20.9908 | -104.4376
20.9946 | -104.431
20.9984 | -104.4231
21.0022 | -104.4161
21.0059 | -104.4111
21.0097 | -104.4054
21.0135 | -104.3972
21.0173 | -104.3891
21.021 | -104.3812
21.0248 | -104.3733
21.0286 | -104.3679
21.0324 | -104.3642
21.0361 | -104.3608
21.0399 | -104.3575
21.0437 | -104.3533
21.0475 | -104.3498
21.0512 | -104.3469
21.055 | -104.3443
21.0588 | -104.3406
21.0626 | -104.3374
21.0663 | -104.3344
21.0701 | -104.3325
21.0739 | -104.3301
21.0777 | -104.3307
21.0814 | -104.3307
21.0852 | -104.3291
21.089 | -104.3281
21.0928 | -104.3262
21.0965 | -104.3256
21.1003 | -104.3262
21.1041 | -104.3274
21.1079 | -104.3294
21.1116 | -104.3312
21.1154 | -104.3349
21.1192 | -104.3377
21.123 | -104.34
21.1267 | -104.3428
21.1305 | -104.3463
21.1343 | -104.3524
21.1381 | -104.3589
21.1418 | -104.365
21.1456 | -104.3707
21.1494 | -104.377
21.1532 | -104.3846
21.1569 | -104.394
21.1607 | -104.4039
21.1645 | -104.4114
21.1683 | -104.4195
21.172 | -104.4306
21.1758 | -104.4417
21.1796 | -104.4535
21.1834 | -104.4651
21.1871 | -104.4777
21.1909 | -104.491
21.1947 | -104.5061
21.1985 | -104.5205
21.2022 | -104.5332
21.206 | -104.5452
21.2098 | -104.5585
21.2136 | -104.5732
21.2173 | -104.5891
21.2211 | -104.6038
21.2249 | -104.617
21.2286 | -104.6302
21.2324 | -104.6455
21.2362 | -104.6587
21.24 | -104.673
21.2437 | -104.6856
21.2475 | -104.6977
21.2513 | -104.7088
21.2551 | -104.7194
21.2588 | -104.7319
21.2626 | -104.7404
21.2664 | -104.7481
21.2702 | -104.7574
21.2739 | -104.7676
21.2777 | -104.7761
21.2815 | -104.783
21.2853 | -104.7901
21.289 | -104.7949
21.2928 | -104.7945
21.2966 | -104.7941
21.3004 | -104.795
21.3041 | -104.7936
21.3079 | -104.7915
21.3117 | -104.7897
21.3155 | -104.7874
21.3192 | -104.7854
21.323 | -104.783
21.3268 | -104.7792
21.3306 | -104.7725
21.3343 | -104.7675
21.3381 | -104.7621
21.3419 | -104.7571
21.3457 | -104.7528
21.3494 | -104.7467
21.3532 | -104.74
21.357 | -104.7347
21.3608 | -104.7294
21.3645 | -104.7225
21.3683 | -104.7178
21.3721 | -104.7114
21.3759 | -104.706
21.3796 | -104.7005
21.3834 | -104.6984
21.3872 | -104.6974
21.391 | -104.6911
21.3947 | -104.6859
21.3985 | -104.6819
21.4023 | -104.6769
21.4061 | -104.6727
21.4098 | -104.6691
21.4136 | -104.6645
21.4174 | -104.6597
21.4212 | -104.6569
21.4249 | -104.6523
21.4287 | -104.6483
21.4325 | -104.6445
21.4363 | -104.6384
21.44 | -104.6338
21.4438 | -104.6279
21.4476 | -104.6255
21.4514 | -104.6205
21.4551 | -104.614
21.4589 | -104.6041
21.4627 | -104.5938
21.4665 | -104.5874
21.4702 | -104.5812
21.474 | -104.574
21.4778 | -104.5657
21.4816 | -104.555
21.4853 | -104.5465
21.4891 | -104.5362
21.4929 | -104.5251
21.4966 | -104.5132
21.5004 | -104.5015
21.5042 | -104.4894
21.508 | -104.4768
21.5117 | -104.4628
21.5155 | -104.4473
21.5193 | -104.4326
21.5231 | -104.4181
21.5268 | -104.4057
21.5306 | -104.3926
21.5344 | -104.3779
21.5382 | -104.3626
21.5419 | -104.3492
21.5457 | -104.3351
21.5495 | -104.325
21.5533 | -104.3144
21.557 | -104.3044
21.5608 | -104.2918
21.5646 | -104.2802
21.5684 | -104.2702
21.5721 | -104.2584
21.5759 | -104.2457
21.5797 | -104.2344
21.5835 | -104.2272
21.5872 | -104.2198
21.591 | -104.211
21.5948 | -104.206
21.5986 | -104.2044
21.6023 | -104.2002
21.6061 | -104.1957
21.6099 | -104.192
21.6137 | -104.1876
21.6174 | -104.185
21.6212 | -104.1805
21.625 | -104.1752
21.6288 | -104.171
21.6325 | -104.1688
21.6363 | -104.1665
21.6401 | -104.1619
21.6439 | -104.1559
21.6476 | -104.1512
21.6514 | -104.147
21.6552 | -104.1432
21.659 | -104.138
21.6627 | -104.1304
21.6665 | -104.1257
21.6703 | -104.1242
21.6741 | -104.1197
21.6778 | -104.1157
21.6816 | -104.1118
21.6854 | -104.1057
21.6892 | -104.0983
21.6929 | -104.0921
21.6967 | -104.0879
21.7005 | -104.0811
21.7043 | -104.0747
21.708 | -104.0679
21.7118 | -104.0602
21.7156 | -104.0544
21.7194 | -104.0492
21.7231 | -104.0412
21.7269 | -104.0328
21.7307 | -104.0252
21.7345 | -104.0162
21.7382 | -104.0085
21.742 | -104.0021
21.7458 | -103.9931
21.7496 | -103.985
21.7533 | -103.9755
21.7571 | -103.9669
21.7609 | -103.9581
21.7646 | -103.9511
21.7684 | -103.9446
21.7722 | -103.9358
21.776 | -103.9258
21.7797 | -103.9143
21.7835 | -103.9054
21.7873 | -103.8989
21.7911 | -103.8922
21.7948 | -103.8842
21.7986 | -103.8776
21.8024 | -103.8707
21.8062 | -103.8643
21.8099 | -103.8576
21.8137 | -103.8518
21.8175 | -103.8448
21.8213 | -103.8389
21.825 | -103.8316
21.8288 | -103.8239
21.8326 | -103.8171
21.8364 | -103.8113
21.8401 | -103.8073
21.8439 | -103.8
21.8477 | -103.7932
21.8515 | -103.7891
21.8552 | -103.7862
21.859 | -103.7821
21.8628 | -103.7805
21.8666 | -103.7784
21.8703 | -103.7764
21.8741 | -103.7737
21.8779 | -103.773
21.8817 | -103.7756
21.8854 | -103.7783
21.8892 | -103.7794
21.893 | -103.7797
21.8968 | -103.781
21.9005 | -103.7832
21.9043 | -103.7843
21.9081 | -103.7868
21.9119 | -103.7904
21.9156 | -103.7947
21.9194 | -103.8016
21.9232 | -103.8084
21.927 | -103.811
21.9307 | -103.8169
21.9345 | -103.8241
21.9383 | -103.8301
21.9421 | -103.8359
21.9458 | -103.8406
21.9496 | -103.8479
21.9534 | -103.858
21.9572 | -103.8663
21.9609 | -103.8731
21.9647 | -103.8805
21.9685 | -103.8888
21.9723 | -103.8967
21.976 | -103.9023
21.9798 | -103.9059
21.9836 | -103.9084
21.9874 | -103.9122
21.9911 | -103.9175
21.9949 | -103.9209
21.9987 | -103.9239
22.0025 | -103.9245
22.0062 | -103.9252
22.01 | -103.9274
22.0138 | -103.9275
22.0176 | -103.9272
22.0213 | -103.9268
22.0251 | -103.9243
22.0289 | -103.9224
22.0327 | -103.9215
22.0364 | -103.9206
22.0402 | -103.9183
22.044 | -103.9157
22.0477 | -103.9125
22.0515 | -103.9096
22.0553 | -103.907
22.0591 | -103.9008
22.0628 | -103.8944
22.0666 | -103.8901
22.0704 | -103.8844
22.0742 | -103.8791
22.0779 | -103.8726
22.0817 | -103.8646
22.0855 | -103.8583
22.0893 | -103.8545
22.093 | -103.849
22.0968 | -103.8402
22.1006 | -103.8334
22.1044 | -103.8275
22.1081 | -103.8194
22.1119 | -103.8108
22.1157 | -103.8053
22.1195 | -103.8012
22.1232 | -103.7974
22.127 | -103.7923
22.1308 | -103.7897
22.1346 | -103.7867
22.1383 | -103.7856
22.1421 | -103.7838
22.1459 | -103.7819
22.1497 | -103.7785
22.1534 | -103.7771
22.1572 | -103.7792
22.161 | -103.7795
22.1648 | -103.7771
22.1685 | -103.777
22.1723 | -103.7783
22.1761 | -103.778
22.1799 | -103.7783
22.1836 | -103.7818
22.1874 | -103.7873
22.1912 | -103.7905
22.195 | -103.7929
22.1987 | -103.7958
22.2025 | -103.8001
22.2063 | -103.8042
22.2101 | -103.8094
22.2138 | -103.8182
22.2176 | -103.8279
22.2214 | -103.8373
22.2252 | -103.8468
22.2289 | -103.856
22.2327 | -103.8639
22.2365 | -103.8728
22.2403 | -103.8848
22.244 | -103.8951
22.2478 | -103.9049
22.2516 | -103.9125
22.2554 | -103.9201
22.2591 | -103.9325
22.2629 | -103.9431
22.2667 | -103.9517
22.2705 | -103.9608
22.2742 | -103.9706
22.278 | -103.9799
22.2818 | -103.9884
22.2856 | -103.9962
22.2893 | -104.0033
22.2931 | -104.0112
22.2969 | -104.0181
22.3007 | -104.0222
22.3044 | -104.0257
22.3082 | -104.0318
22.312 | -104.0368
22.3158 | -104.0418
22.3195 | -104.0453
22.3233 | -104.0475
22.3271 | -104.0441
22.3308 | -104.0453
22.3346 | -104.0463
22.3384 | -104.0463
22.3422 | -104.0457
22.3459 | -104.0443
22.3497 | -104.0442
22.3535 | -104.0426
22.3573 | -104.0402
22.361 | -104.0384
22.3648 | -104.037
22.3686 | -104.0332
22.3724 | -104.0312
22.3761 | -104.0306
22.3799 | -104.0298
22.3837 | -104.0255
22.3875 | -104.0224
22.3912 | -104.0203
22.395 | -104.0186
22.3988 | -104.0199
22.4026 | -104.0197
22.4063 | -104.0184
22.4101 | -104.0187
22.4139 | -104.0197
22.4177 | -104.0194
22.4214 | -104.0198
22.4252 | -104.0225
22.429 | -104.025
22.4328 | -104.0257
22.4365 | -104.0269
22.4403 | -104.0279
22.4441 | -104.0312
22.4479 | -104.0326
22.4516 | -104.0358
22.4554 | -104.0395
22.4592 | -104.043
22.463 | -104.045
22.4667 | -104.049
22.4705 | -104.0514
22.4743 | -104.0522
22.4781 | -104.0565
22.4818 | -104.0603
22.4856 | -104.0608
22.4894 | -104.0609
22.4932 | -104.0629
22.4969 | -104.0636
22.5007 | -104.0625
22.5045 | -104.0619
22.5083 | -104.062
22.512 | -104.0608
22.5158 | -104.0578
22.5196 | -104.0545
22.5234 | -104.0503
22.5271 | -104.0422
22.5309 | -104.0333
22.5347 | -104.0269
22.5385 | -104.0177
22.5422 | -104.0066
22.546 | -103.9962
22.5498 | -103.9865
22.5536 | -103.9759
22.5573 | -103.9642
22.5611 | -103.953
22.5649 | -103.9406
22.5687 | -103.9299
22.5724 | -103.9199
22.5762 | -103.9053
22.58 | -103.8896
22.5837 | -103.8749
22.5875 | -103.8611
22.5913 | -103.8461
22.5951 | -103.8309
22.5988 | -103.8163
22.6026 | -103.8041
22.6064 | -103.7924
22.6102 | -103.7785
22.6139 | -103.7647
22.6177 | -103.7527
22.6215 | -103.7377
22.6253 | -103.723
22.629 | -103.7103
22.6328 | -103.699
22.6366 | -103.6883
22.6404 | -103.6767
22.6441 | -103.6659
22.6479 | -103.6572
22.6517 | -103.649
22.6555 | -103.6392
22.6592 | -103.6288
22.663 | -103.6223
22.6668 | -103.6171
22.6706 | -103.6088
22.6743 | -103.6017
22.6781 | -103.5978
22.6819 | -103.5946
22.6857 | -103.5918
22.6894 | -103.5913
22.6932 | -103.5891
22.697 | -103.5876
22.7008 | -103.5885
22.7045 | -103.5912
22.7083 | -103.596
22.7121 | -103.6002
22.7159 | -103.6043
22.7196 | -103.6069
22.7234 | -103.6112
22.7272 | -103.6137
22.731 | -103.6162
22.7347 | -103.6198
22.7385 | -103.6242
22.7423 | -103.6269
22.7461 | -103.6308
22.7498 | -103.6377
22.7536 | -103.6444
22.7574 | -103.6491
22.7612 | -103.6565
22.7649 | -103.6631
22.7687 | -103.6728
22.7725 | -103.6799
22.7763 | -103.6862
22.78 | -103.6934
22.7838 | -103.7013
22.7876 | -103.7055
22.7914 | -103.7115
22.7951 | -103.7199
22.7989 | -103.7251
22.8027 | -103.7305
22.8065 | -103.7363
22.8102 | -103.743
22.814 | -103.7485
22.8178 | -103.7547
22.8216 | -103.7599
22.8253 | -103.7619
22.8291 | -103.7632
22.8329 | -103.7663
22.8367 | -103.7707
22.8404 | -103.7721
22.8442 | -103.7742
22.848 | -103.7786
22.8517 | -103.7796
22.8555 | -103.7779
22.8593 | -103.7783
22.8631 | -103.7775
22.8668 | -103.7759
22.8706 | -103.7717
22.8744 | -103.7687
22.8782 | -103.765
22.8819 | -103.7613
22.8857 | -103.7554
22.8895 | -103.751
22.8933 | -103.7467
22.897 | -103.7411
22.9008 | -103.7365
22.9046 | -103.7326
22.9084 | -103.7269
22.9121 | -103.7202
22.9159 | -103.7132
22.9197 | -103.7058
22.9235 | -103.6976
22.9272 | -103.6872
22.931 | -103.6794
22.9348 | -103.6725
22.9386 | -103.6659
22.9423 | -103.6579
22.9461 | -103.6513
22.9499 | -103.6425
22.9537 | -103.6322
22.9574 | -103.6237
22.9612 | -103.617
22.965 | -103.6097
22.9688 | -103.6021
22.9725 | -103.5952
22.9763 | -103.5896
22.9801 | -103.5872
22.9839 | -103.5849
22.9876 | -103.5783
22.9914 | -103.5737
22.9952 | -103.5721
22.999 | -103.5704
23.0027 | -103.5687
23.0065 | -103.5668
23.0103 | -103.5643
23.0141 | -103.563
23.0178 | -103.5652
23.0216 | -103.5687
23.0254 | -103.5708
23.0292 | -103.5739
23.0329 | -103.5767
23.0367 | -103.5803
23.0405 | -103.5848
23.0443 | -103.592
23.048 | -103.5997
23.0518 | -103.6089
23.0556 | -103.6184
23.0594 | -103.6291
23.0631 | -103.6376
23.0669 | -103.6468
23.0707 | -103.6548
23.0745 | -103.6629
23.0782 | -103.671
23.082 | -103.6796
23.0858 | -103.6869
23.0896 | -103.6966
23.0933 | -103.7062
23.0971 | -103.7165
23.1009 | -103.7248
23.1047 | -103.7336
23.1084 | -103.7443
23.1122 | -103.7558
23.116 | -103.7647
23.1198 | -103.7724
23.1235 | -103.7803
23.1273 | -103.7875
23.1311 | -103.7941
23.1348 | -103.8005
23.1386 | -103.8085
23.1424 | -103.8164
23.1462 | -103.8247
23.1499 | -103.8335
23.1537 | -103.8415
23.1575 | -103.8478
23.1613 | -103.8524
23.165 | -103.859
23.1688 | -103.865
23.1726 | -103.8719
23.1764 | -103.8787
23.1801 | -103.8854
23.1839 | -103.8926
23.1877 | -103.9012
23.1915 | -103.9068
23.1952 | -103.911
23.199 | -103.9159
23.2028 | -103.9227
23.2066 | -103.9308
23.2103 | -103.9376
23.2141 | -103.9454
23.2179 | -103.9504
23.2217 | -103.9539
23.2254 | -103.9589
23.2292 | -103.9653
23.233 | -103.9713
23.2368 | -103.9782
23.2405 | -103.9846
23.2443 | -103.9923
23.2481 | -104.002
23.2519 | -104.0117
23.2556 | -104.0229
23.2594 | -104.0313
23.2632 | -104.039
23.267 | -104.0478
23.2707 | -104.0578
23.2745 | -104.0682
23.2783 | -104.0799
23.2821 | -104.0922
23.2858 | -104.1025
23.2896 | -104.1112
23.2934 | -104.1218
23.2972 | -104.1317
23.3009 | -104.1413
23.3047 | -104.1504
23.3085 | -104.157
23.3123 | -104.1654
23.316 | -104.1758
23.3198 | -104.1862
23.3236 | -104.1957
23.3274 | -104.205
23.3311 | -104.2138
23.3349 | -104.2227
23.3387 | -104.2324
23.3425 | -104.2424
23.3462 | -104.2509
23.35 | -104.258
23.3538 | -104.2648
23.3576 | -104.2722
23.3613 | -104.2777
23.3651 | -104.2844
23.3689 | -104.2906
23.3727 | -104.2957
23.3764 | -104.3018
23.3802 | -104.3098
23.384 | -104.3164
23.3878 | -104.3219
23.3915 | -104.3293
23.3953 | -104.3375
23.3991 | -104.3448
23.4029 | -104.352
23.4066 | -104.3596
23.4104 | -104.3668
23.4142 | -104.3748
23.4179 | -104.3827
23.4217 | -104.3899
23.4255 | -104.3948
23.4293 | -104.4009
23.433 | -104.4107
23.4368 | -104.4207
23.4406 | -104.4286
23.4444 | -104.4376
23.4481 | -104.4474
23.4519 | -104.4573
23.4557 | -104.4674
23.4595 | -104.4789
23.4632 | -104.4908
23.467 | -104.5015
23.4708 | -104.5112
23.4746 | -104.5198
23.4783 | -104.5289
23.4821 | -104.5388
23.4859 | -104.5503
23.4897 | -104.5612
23.4934 | -104.5731
23.4972 | -104.5852
23.501 | -104.5955
23.5048 | -104.6028
23.5085 | -104.6109
23.5123 | -104.6196
23.5161 | -104.6271
23.5199 | -104.6349
23.5236 | -104.6446
23.5274 | -104.6527
23.5312 | -104.6572
23.535 | -104.6618
23.5387 | -104.667
23.5425 | -104.6733
23.5463 | -104.678
23.5501 | -104.6829
23.5538 | -104.6868
23.5576 | -104.6892
23.5614 | -104.6898
23.5652 | -104.689
23.5689 | -104.6872
23.5727 | -104.685
23.5765 | -104.6821
23.5803 | -104.6769
23.584 | -104.6714
23.5878 | -104.6669
23.5916 | -104.6612
23.5954 | -104.6555
23.5991 | -104.6513
23.6029 | -104.647
23.6067 | -104.6421
23.6105 | -104.6355
23.6142 | -104.6254
23.618 | -104.615
23.6218 | -104.608
23.6256 | -104.6022
23.6293 | -104.5977
23.6331 | -104.5902
23.6369 | -104.5816
23.6407 | -104.5765
23.6444 | -104.5708
23.6482 | -104.5647
23.652 | -104.5573
23.6558 | -104.5507
23.6595 | -104.5447
23.6633 | -104.5379
23.6671 | -104.5315
23.6709 | -104.5245
23.6746 | -104.5168
23.6784 | -104.5102
23.6822 | -104.5047
23.6859 | -104.5009
23.6897 | -104.4978
23.6935 | -104.4953
23.6973 | -104.4919
23.701 | -104.4894
23.7048 | -104.4872
23.7086 | -104.4845
23.7124 | -104.4821
23.7161 | -104.4823
23.7199 | -104.4834
23.7237 | -104.4845
23.7275 | -104.4864
23.7312 | -104.487
23.735 | -104.4865
23.7388 | -104.4874
23.7426 | -104.4912
23.7463 | -104.4963
23.7501 | -104.5004
23.7539 | -104.5039
23.7577 | -104.5082
23.7614 | -104.5151
23.7652 | -104.5214
23.769 | -104.5284
23.7728 | -104.5342
23.7765 | -104.5425
23.7803 | -104.5495
23.7841 | -104.5566
23.7879 | -104.5636
23.7916 | -104.5732
23.7954 | -104.5811
23.7992 | -104.5862
23.803 | -104.5885
23.8067 | -104.5948
23.8105 | -104.6024
23.8143 | -104.6094
23.8181 | -104.616
23.8218 | -104.6214
23.8256 | -104.6256
23.8294 | -104.6287
23.8332 | -104.6331
23.8369 | -104.6387
23.8407 | -104.6448
23.8445 | -104.6489
23.8483 | -104.6511
23.852 | -104.6524
23.8558 | -104.6546
23.8596 | -104.6577
23.8634 | -104.6603
23.8671 | -104.6614
23.8709 | -104.6594
23.8747 | -104.6568
23.8785 | -104.6547
23.8822 | -104.6526
23.886 | -104.6496
23.8898 | -104.6464
23.8936 | -104.6422
23.8973 | -104.6375
23.9011 | -104.6315
23.9049 | -104.625
23.9087 | -104.618
23.9124 | -104.6105
23.9162 | -104.6035
23.92 | -104.5978
23.9238 | -104.5893
23.9275 | -104.5807
23.9313 | -104.5716
23.9351 | -104.5614
23.9389 | -104.552
23.9426 | -104.5424
23.9464 | -104.5345
23.9502 | -104.5294
23.9539 | -104.5245
23.9577 | -104.5199
23.9615 | -104.5129
23.9653 | -104.5049
23.969 | -104.4974
23.9728 | -104.4914
23.9766 | -104.4871
23.9804 | -104.4823
23.9841 | -104.4775
23.9879 | -104.4733
23.9917 | -104.47
23.9955 | -104.4698
23.9992 | -104.4707
24.003 | -104.4725
24.0068 | -104.4736
24.0106 | -104.476
24.0143 | -104.4797
24.0181 | -104.4819
24.0219 | -104.4851
24.0257 | -104.4878
24.0294 | -104.4915
24.0332 | -104.498
24.037 | -104.5044
24.0408 | -104.5101
24.0445 | -104.516
24.0483 | -104.5239
24.0521 | -104.5316
24.0559 | -104.5383
24.0596 | -104.5452
24.0634 | -104.5538
24.0672 | -104.5639
24.071 | -104.5745
24.0747 | -104.5845
24.0785 | -104.5927
24.0823 | -104.6018
24.0861 | -104.6121
24.0898 | -104.622
24.0936 | -104.633
24.0974 | -104.6426
24.1012 | -104.6508
24.1049 | -104.6593
24.1087 | -104.6671
24.1125 | -104.6767
24.1163 | -104.6856
24.12 | -104.6943
24.1238 | -104.704
24.1276 | -104.7135
24.1314 | -104.7203
24.1351 | -104.7273
24.1389 | -104.7377
24.1427 | -104.7485
24.1465 | -104.7549
24.1502 | -104.7612
24.154 | -104.7678
24.1578 | -104.774
24.1616 | -104.7794
24.1653 | -104.7846
24.1691 | -104.7901
24.1729 | -104.7931
24.1767 | -104.7958
24.1804 | -104.7989
24.1842 | -104.8032
24.188 | -104.8068
24.1918 | -104.81
24.1955 | -104.8109
24.1993 | -104.8099
24.2031 | -104.8104
24.2069 | -104.8118
24.2106 | -104.8146
24.2144 | -104.8145
24.2182 | -104.8127
24.2219 | -104.8112
24.2257 | -104.8107
24.2295 | -104.8113
24.2333 | -104.8106
24.237 | -104.8096
24.2408 | -104.8104
24.2446 | -104.8108
24.2484 | -104.8104
24.2521 | -104.81
24.2559 | -104.8077
24.2597 | -104.8042
24.2635 | -104.802
24.2672 | -104.8007
24.271 | -104.7992
24.2748 | -104.7962
24.2786 | -104.7925
24.2823 | -104.7871
24.2861 | -104.7794
24.2899 | -104.7701
24.2937 | -104.7597
24.2974 | -104.7503
24.3012 | -104.7392
24.305 | -104.7266
24.3088 | -104.7135
24.3125 | -104.7026
24.3163 | -104.6917
24.3201 | -104.679
24.3239 | -104.6674
24.3276 | -104.654
24.3314 | -104.6371
24.3352 | -104.619
24.339 | -104.598
24.3427 | -104.5769
24.3465 | -104.5564
24.3503 | -104.5367
24.3541 | -104.5148
24.3578 | -104.4894
24.3616 | -104.4662
24.3654 | -104.4426
24.3692 | -104.4209
24.3729 | -104.3974
24.3767 | -104.3718
24.3805 | -104.3456
24.3843 | -104.3208
24.388 | -104.2954
24.3918 | -104.2686
24.3956 | -104.2417
24.3994 | -104.2133
24.4031 | -104.1854
24.4069 | -104.1569
24.4107 | -104.1285
24.4145 | -104.1024
24.4182 | -104.0783
24.422 | -104.0559
24.4258 | -104.0328
24.4296 | -104.0092
24.4333 | -103.9867
24.4371 | -103.9654
24.4409 | -103.9459
24.4447 | -103.927
24.4484 | -103.9099
24.4522 | -103.8936
24.456 | -103.8772
24.4598 | -103.8632
24.4635 | -103.8509
24.4673 | -103.8393
24.4711 | -103.8299
24.4749 | -103.8224
24.4786 | -103.8162
24.4824 | -103.8101
24.4862 | -103.806
24.49 | -103.8021
24.4937 | -103.8001
24.4975 | -103.8013
24.5013 | -103.8021
24.505 | -103.8157
24.5088 | -103.8195
24.5126 | -103.8238
24.5164 | -103.8287
24.5201 | -103.8357
24.5239 | -103.8427
24.5277 | -103.8507
24.5315 | -103.8599
24.5352 | -103.871
24.539 | -103.884
24.5428 | -103.8982
24.5466 | -103.9123
24.5503 | -103.9242
24.5541 | -103.9379
24.5579 | -103.9522
24.5617 | -103.967
24.5654 | -103.9799
24.5692 | -103.9913
24.573 | -103.9985
24.5768 | -104.0062
24.5805 | -104.0125
24.5843 | -104.0192
24.5881 | -104.0254
24.5919 | -104.0325
24.5956 | -104.041
24.5994 | -104.0454
24.6032 | -104.049
24.607 | -104.0541
24.6107 | -104.0584
24.6145 | -104.0606
24.6183 | -104.0613
24.6221 | -104.0618
24.6258 | -104.0626
24.6296 | -104.063
24.6334 | -104.0635
24.6372 | -104.0625
24.6409 | -104.0611
24.6447 | -104.0563
24.6485 | -104.051
24.6523 | -104.0462
24.656 | -104.0387
24.6598 | -104.0276
24.6636 | -104.0157
24.6674 | -104.0024
24.6711 | -103.9886
24.6749 | -103.9762
24.6787 | -103.9654
24.6825 | -103.9543
24.6862 | -103.9433
24.69 | -103.933
24.6938 | -103.9225
24.6976 | -103.9106
24.7013 | -103.899
24.7051 | -103.8892
24.7089 | -103.879
24.7127 | -103.8694
24.7164 | -103.8579
24.7202 | -103.8466
24.724 | -103.8374
24.7278 | -103.8283
24.7315 | -103.819
24.7353 | -103.8117
24.7391 | -103.8055
24.7429 | -103.8003
24.7466 | -103.797
24.7504 | -103.795
24.7542 | -103.7937
24.758 | -103.7919
24.7617 | -103.7908
24.7655 | -103.79
24.7693 | -103.7883
24.773 | -103.7878
24.7768 | -103.7903
24.7806 | -103.7958
24.7844 | -103.8008
24.7881 | -103.806
24.7919 | -103.8122
24.7957 | -103.8197
24.7995 | -103.8271
24.8032 | -103.8355
24.807 | -103.8434
24.8108 | -103.853
24.8146 | -103.8628
24.8183 | -103.8717
24.8221 | -103.8811
24.8259 | -103.8909
24.8297 | -103.9031
24.8334 | -103.9152
24.8372 | -103.9266
24.841 | -103.9356
24.8448 | -103.9409
24.8485 | -103.9505
24.8523 | -103.9602
24.8561 | -103.9697
24.8599 | -103.9779
24.8636 | -103.9863
24.8674 | -103.9956
24.8712 | -104.0044
24.875 | -104.0126
24.8787 | -104.0203
24.8825 | -104.0276
24.8863 | -104.033
24.8901 | -104.0358
24.8938 | -104.0401
24.8976 | -104.044
24.9014 | -104.0483
24.9052 | -104.0489
24.9089 | -104.0495
24.9127 | -104.0503
24.9165 | -104.0525
24.9203 | -104.0548
24.924 | -104.059
24.9278 | -104.0611
24.9316 | -104.0633
24.9354 | -104.0637
24.9391 | -104.0626
24.9429 | -104.0578
24.9467 | -104.0521
24.9505 | -104.0476
24.9542 | -104.0405
24.958 | -104.0307
24.9618 | -104.0233
24.9656 | -104.0177
24.9693 | -104.0089
24.9731 | -103.998
24.9769 | -103.9876
24.9807 | -103.9802
24.9844 | -103.9733
24.9882 | -103.9664
24.992 | -103.9584
24.9958 | -103.9529
24.9995 | -103.9484
25.0033 | -103.9427
25.0071 | -103.9387
25.0109 | -103.9342
25.0146 | -103.9296
25.0184 | -103.9262
25.0222 | -103.9206
25.026 | -103.9148
25.0297 | -103.9108
25.0335 | -103.9072
25.0373 | -103.904
25.041 | -103.9007
25.0448 | -103.8972
25.0486 | -103.9014
25.0524 | -103.9002
25.0561 | -103.9019
25.0599 | -103.903
25.0637 | -103.9032
25.0675 | -103.9038
25.0712 | -103.9044
25.075 | -103.9062
25.0788 | -103.908
25.0826 | -103.9121
25.0863 | -103.9173
25.0901 | -103.9218
25.0939 | -103.9278
25.0977 | -103.9339
25.1014 | -103.9385
25.1052 | -103.942
25.109 | -103.9465
25.1128 | -103.9537
25.1165 | -103.961
25.1203 | -103.9692
25.1241 | -103.9781
25.1279 | -103.9848
25.1316 | -103.9912
25.1354 | -103.9995
25.1392 | -104.0075
25.143 | -104.0161
25.1467 | -104.0266
25.1505 | -104.0349
25.1543 | -104.0434
25.1581 | -104.0501
25.1618 | -104.0571
25.1656 | -104.064
25.1694 | -104.0714
25.1732 | -104.0801
25.1769 | -104.0893
25.1807 | -104.0977
25.1845 | -104.1042
25.1883 | -104.1114
25.192 | -104.1192
25.1958 | -104.1276
25.1996 | -104.1374
25.2034 | -104.142
25.2071 | -104.1455
25.2109 | -104.1492
25.2147 | -104.1523
25.2185 | -104.1563
25.2222 | -104.1575
25.226 | -104.158
25.2298 | -104.1591
25.2336 | -104.159
25.2373 | -104.1573
25.2411 | -104.157
25.2449 | -104.1572
25.2487 | -104.1531
25.2524 | -104.1486
25.2562 | -104.1453
25.26 | -104.1439
25.2638 | -104.1409
25.2675 | -104.1365
25.2713 | -104.1318
25.2751 | -104.1266
25.2789 | -104.1214
25.2826 | -104.1149
25.2864 | -104.1078
25.2902 | -104.0998
25.294 | -104.091
25.2977 | -104.0836
25.3015 | -104.0764
25.3053 | -104.0691
25.309 | -104.0621
25.3128 | -104.0561
25.3166 | -104.0472
25.3204 | -104.0371
25.3241 | -104.0322
25.3279 | -104.0271
25.3317 | -104.0197
25.3355 | -104.0099
25.3392 | -103.9987
25.343 | -103.9895
25.3468 | -103.9819
25.3506 | -103.9746
25.3543 | -103.967
25.3581 | -103.9574
25.3619 | -103.948
25.3657 | -103.9409
25.3694 | -103.9343
25.3732 | -103.9279
25.377 | -103.923
25.3808 | -103.9184
25.3845 | -103.9115
25.3883 | -103.9058
25.3921 | -103.9011
25.3959 | -103.8943
25.3996 | -103.8872
25.4034 | -103.8796
25.4072 | -103.8741
25.411 | -103.8707
25.4147 | -103.8665
25.4185 | -103.8641
25.4223 | -103.8632
25.4261 | -103.8623
25.4298 | -103.862
25.4336 | -103.8632
25.4374 | -103.863
25.4412 | -103.8619
25.4449 | -103.8606
25.4487 | -103.8594
25.4525 | -103.8575
25.4563 | -103.8577
25.46 | -103.8589
25.4638 | -103.862
25.4676 | -103.8644
25.4714 | -103.8649
25.4751 | -103.8659
25.4789 | -103.8676
25.4827 | -103.8696
25.4865 | -103.8724
25.4902 | -103.8746
25.494 | -103.8743
25.4978 | -103.8766
25.5016 | -103.8786
25.5053 | -103.879
25.5091 | -103.8813
25.5129 | -103.8849
25.5167 | -103.8891
25.5204 | -103.8923
25.5242 | -103.8945
25.528 | -103.8968
25.5318 | -103.8998
25.5355 | -103.9021
25.5393 | -103.9044
25.5431 | -103.9062
25.5469 | -103.9074
25.5506 | -103.9087
25.5544 | -103.9102
25.5582 | -103.9117
25.562 | -103.9141
25.5657 | -103.9168
25.5695 | -103.9181
25.5733 | -103.9204
25.577 | -103.9213
25.5808 | -103.9239
25.5846 | -103.9262
25.5884 | -103.9258
25.5921 | -103.9264
25.5959 | -103.9282
25.5997 | -103.931
25.6035 | -103.9323
25.6072 | -103.9342
25.611 | -103.937
25.6148 | -103.9409
25.6186 | -103.944
25.6223 | -103.9465
25.6261 | -103.9485
25.6299 | -103.9523
25.6337 | -103.9543
25.6374 | -103.9559
25.6412 | -103.958
25.645 | -103.9598
25.6488 | -103.9621
25.6525 | -103.965
25.6563 | -103.9688
25.6601 | -103.9733
25.6639 | -103.9765
25.6676 | -103.9795
25.6714 | -103.9828
25.6752 | -103.985
25.679 | -103.9882
25.6827 | -103.9923
25.6865 | -103.9937
25.6903 | -103.9967
25.6941 | -103.9995
25.6978 | -104.0011
25.7016 | -104.0037
25.7054 | -104.0058
25.7092 | -104.0075
25.7129 | -104.0087
25.7167 | -104.0104
25.7205 | -104.0141
25.7243 | -104.0188
25.728 | -104.0228
25.7318 | -104.0255
25.7356 | -104.0278
25.7394 | -104.0306
25.7431 | -104.0329
25.7469 | -104.0368
25.7507 | -104.0395
25.7545 | -104.0425
25.7582 | -104.0486
25.762 | -104.0552
25.7658 | -104.0605
25.7696 | -104.0641
25.7733 | -104.0673
25.7771 | -104.07
25.7809 | -104.0725
25.7847 | -104.0763
25.7884 | -104.0797
25.7922 | -104.0828
25.796 | -104.0855
25.7998 | -104.0907
25.8035 | -104.0963
25.8073 | -104.1015
25.8111 | -104.1068
25.8149 | -104.1125
25.8186 | -104.1174
25.8224 | -104.1224
25.8262 | -104.1264
25.83 | -104.1294
25.8337 | -104.1347
25.8375 | -104.14
25.8413 | -104.1449
25.8451 | -104.1503
25.8488 | -104.1569
25.8526 | -104.1635
25.8564 | -104.1694
25.8601 | -104.1744
25.8639 | -104.181
25.8677 | -104.1878
25.8715 | -104.1932
25.8752 | -104.1986
25.879 | -104.2048
25.8828 | -104.212
25.8866 | -104.2206
25.8903 | -104.2276
25.8941 | -104.231
25.8979 | -104.2371
25.9017 | -104.2454
25.9054 | -104.2536
25.9092 | -104.2609
25.913 | -104.2696
25.9168 | -104.2786
25.9205 | -104.2869
25.9243 | -104.2958
25.9281 | -104.3032
25.9319 | -104.3118
25.9356 | -104.3214
25.9394 | -104.3309
25.9432 | -104.3382
25.947 | -104.3467
25.9507 | -104.3566
25.9545 | -104.3627
25.9583 | -104.3689
25.9621 | -104.3758
25.9658 | -104.3826
25.9696 | -104.3887
25.9734 | -104.3949
25.9772 | -104.4026
25.9809 | -104.4086
25.9847 | -104.4127
25.9885 | -104.4159
25.9923 | -104.4198
25.996 | -104.4232
25.9998 | -104.4286
26.0036 | -104.4338
26.0074 | -104.4365
26.0111 | -104.4412
26.0149 | -104.4435
26.0187 | -104.4454
26.0225 | -104.446
26.0262 | -104.4457
26.03 | -104.445
26.0338 | -104.4455
26.0376 | -104.4459
26.0413 | -104.4452
26.0451 | -104.4441
26.0489 | -104.4429
26.0527 | -104.4431
26.0564 | -104.4417
26.0602 | -104.4402
26.064 | -104.4408
26.0678 | -104.4408
26.0715 | -104.438
26.0753 | -104.4344
26.0791 | -104.4317
26.0829 | -104.4294
26.0866 | -104.4273
26.0904 | -104.4245
26.0942 | -104.422
26.098 | -104.4205
26.1017 | -104.4204
26.1055 | -104.4193
26.1093 | -104.418
26.1131 | -104.416
26.1168 | -104.4147
26.1206 | -104.4132
26.1244 | -104.411
26.1281 | -104.41
26.1319 | -104.4098
26.1357 | -104.4121
26.1395 | -104.4142
26.1432 | -104.4148
26.147 | -104.4171
26.1508 | -104.4203
26.1546 | -104.4211
26.1583 | -104.4225
26.1621 | -104.4248
26.1659 | -104.4281
26.1697 | -104.4315
26.1734 | -104.4358
26.1772 | -104.4395
26.181 | -104.4435
26.1848 | -104.449
26.1885 | -104.4561
26.1923 | -104.4635
26.1961 | -104.4688
26.1999 | -104.476
26.2036 | -104.481
26.2074 | -104.4862
26.2112 | -104.4931
26.215 | -104.501
26.2187 | -104.5079
26.2225 | -104.5134
26.2263 | -104.5203
26.2301 | -104.527
26.2338 | -104.5334
26.2376 | -104.54
26.2414 | -104.5484
26.2452 | -104.5558
26.2489 | -104.5622
26.2527 | -104.5673
26.2565 | -104.5721
26.2603 | -104.5777
26.264 | -104.5855
26.2678 | -104.5893
26.2716 | -104.5903
26.2754 | -104.5923
26.2791 | -104.5954
26.2829 | -104.5966
26.2867 | -104.596
26.2905 | -104.5957
26.2942 | -104.5936
26.298 | -104.5922
26.3018 | -104.5907
26.3056 | -104.586
26.3093 | -104.5806
26.3131 | -104.5767
26.3169 | -104.5722
26.3207 | -104.5683
26.3244 | -104.5655
26.3282 | -104.5603
26.332 | -104.5534
26.3358 | -104.546
26.3395 | -104.5372
26.3433 | -104.5265
26.3471 | -104.5162
26.3509 | -104.5064
26.3546 | -104.4979
26.3584 | -104.4915
26.3622 | -104.4813
26.366 | -104.47
26.3697 | -104.4588
26.3735 | -104.4494
26.3773 | -104.441
26.3811 | -104.4317
26.3848 | -104.4243
26.3886 | -104.4152
26.3924 | -104.4053
26.3961 | -104.3961
26.3999 | -104.3887
26.4037 | -104.3808
26.4075 | -104.374
26.4112 | -104.3677
26.415 | -104.3604
26.4188 | -104.354
26.4226 | -104.3479
26.4263 | -104.3422
26.4301 | -104.3374
26.4339 | -104.3331
26.4377 | -104.3281
26.4414 | -104.3247
26.4452 | -104.3226
26.449 | -104.3207
26.4528 | -104.318
26.4565 | -104.3154
26.4603 | -104.3124
26.4641 | -104.3097
26.4679 | -104.3072
26.4716 | -104.3022
26.4754 | -104.2995
26.4792 | -104.2985
26.483 | -104.2967
26.4867 | -104.2937
26.4905 | -104.2918
26.4943 | -104.2913
26.4981 | -104.2904
26.5018 | -104.2896
26.5056 | -104.2903
26.5094 | -104.2902
26.5132 | -104.2883
26.5169 | -104.2865
26.5207 | -104.2854
26.5245 | -104.284
26.5283 | -104.2837
26.532 | -104.2831
26.5358 | -104.283
26.5396 | -104.2816
26.5434 | -104.2786
26.5471 | -104.2763
26.5509 | -104.2765
26.5547 | -104.2764
26.5585 | -104.2736
26.5622 | -104.2713
26.566 | -104.2706
26.5698 | -104.2718
26.5736 | -104.2727
26.5773 | -104.2725
26.5811 | -104.2724
26.5849 | -104.2705
26.5887 | -104.2705
26.5924 | -104.2712
26.5962 | -104.2715
26.6 | -104.2719
26.6038 | -104.272
26.6075 | -104.2726
26.6113 | -104.274
26.6151 | -104.275
26.6189 | -104.2764
26.6226 | -104.277
26.6264 | -104.2776
26.6302 | -104.2793
26.634 | -104.2813
26.6377 | -104.2822
26.6415 | -104.2839
26.6453 | -104.287
26.6491 | -104.2897
26.6528 | -104.2912
26.6566 | -104.2929
26.6604 | -104.2966
26.6641 | -104.2997
26.6679 | -104.304
26.6717 | -104.3072
26.6755 | -104.3098
26.6792 | -104.3151
26.683 | -104.3217
26.6868 | -104.3257
26.6906 | -104.3291
26.6943 | -104.3318
26.6981 | -104.3347
26.7019 | -104.3395
26.7057 | -104.3442
26.7094 | -104.3478
26.7132 | -104.3526
26.717 | -104.3597
26.7208 | -104.3662
26.7245 | -104.3717
26.7283 | -104.379
26.7321 | -104.386
26.7359 | -104.3904
26.7396 | -104.3967
26.7434 | -104.4037
26.7472 | -104.4115
26.751 | -104.419
26.7547 | -104.4258
26.7585 | -104.4329
26.7623 | -104.4399
26.7661 | -104.4476
26.7698 | -104.4559
26.7736 | -104.463
26.7774 | -104.4699
26.7812 | -104.4766
26.7849 | -104.483
26.7887 | -104.4893
26.7925 | -104.4973
26.7963 | -104.5043
26.8 | -104.5109
26.8038 | -104.5172
26.8076 | -104.5223
26.8114 | -104.5288
26.8151 | -104.537
26.8189 | -104.5435
26.8227 | -104.5486
26.8265 | -104.5541
26.8302 | -104.5599
26.834 | -104.5662
26.8378 | -104.5733
26.8416 | -104.5798
26.8453 | -104.5846
26.8491 | -104.5913
26.8529 | -104.5943
26.8567 | -104.5977
26.8604 | -104.6039
26.8642 | -104.6104
26.868 | -104.6135
26.8718 | -104.6142
26.8755 | -104.6158
26.8793 | -104.6183
26.8831 | -104.6204
26.8869 | -104.6221
26.8906 | -104.6255
26.8944 | -104.6291
26.8982 | -104.6307
26.902 | -104.6315
26.9057 | -104.6325
26.9095 | -104.6345
26.9133 | -104.637
26.9171 | -104.6388
26.9208 | -104.6412
26.9246 | -104.6395
26.9284 | -104.6367
26.9322 | -104.6343
26.9359 | -104.6334
26.9397 | -104.6313
26.9435 | -104.6284
26.9472 | -104.6252
26.951 | -104.6222
26.9548 | -104.6201
26.9586 | -104.6163
26.9623 | -104.6137
26.9661 | -104.6086
26.9699 | -104.6022
26.9737 | -104.5975
26.9774 | -104.5899
26.9812 | -104.5833
26.985 | -104.5763
26.9888 | -104.5672
26.9925 | -104.5567
26.9963 | -104.5468
27.0001 | -104.5381
27.0039 | -104.5303
27.0076 | -104.5206
27.0114 | -104.5119
27.0152 | -104.5016
27.019 | -104.4901
27.0227 | -104.478
27.0265 | -104.4655
27.0303 | -104.4556
27.0341 | -104.445
27.0378 | -104.4326
27.0416 | -104.42
27.0454 | -104.407
27.0492 | -104.3952
27.0529 | -104.3823
27.0567 | -104.3681
27.0605 | -104.3534
27.0643 | -104.34
27.068 | -104.328
27.0718 | -104.3152
27.0756 | -104.3013
27.0794 | -104.2884
27.0831 | -104.2747
27.0869 | -104.2616
27.0907 | -104.2497
27.0945 | -104.2362
27.0982 | -104.2212
27.102 | -104.2082
27.1058 | -104.1944
27.1096 | -104.1804
27.1133 | -104.1672
27.1171 | -104.1548
27.1209 | -104.1427
27.1247 | -104.1302
27.1284 | -104.119
27.1322 | -104.1069
27.136 | -104.0961
27.1398 | -104.0866
27.1435 | -104.0761
27.1473 | -104.067
27.1511 | -104.0591
27.1549 | -104.0497
27.1586 | -104.0403
27.1624 | -104.0334
27.1662 | -104.0263
27.17 | -104.0207
27.1737 | -104.0133
27.1775 | -104.0071
27.1813 | -104.0028
27.1851 | -103.9983
27.1888 | -103.9949
27.1926 | -103.9918
27.1964 | -103.9883
27.2002 | -103.985
27.2039 | -103.9829
27.2077 | -103.9808
27.2115 | -103.9811
27.2153 | -103.9825
27.219 | -103.9821
27.2228 | -103.9822
27.2266 | -103.9895
27.2303 | -103.9938
27.2341 | -103.9976
27.2379 | -104.0016
27.2417 | -104.0051
27.2454 | -104.0078
27.2492 | -104.0123
27.253 | -104.0193
27.2568 | -104.0254
27.2605 | -104.0299
27.2643 | -104.0346
27.2681 | -104.0386
27.2719 | -104.0443
27.2756 | -104.0494
27.2794 | -104.0536
27.2832 | -104.0584
27.287 | -104.0652
27.2907 | -104.07
27.2945 | -104.074
27.2983 | -104.0791
27.3021 | -104.0848
27.3058 | -104.0902
27.3096 | -104.096
27.3134 | -104.0991
27.3172 | -104.1022
27.3209 | -104.1058
27.3247 | -104.1088
27.3285 | -104.1105
27.3323 | -104.1134
27.336 | -104.1143
27.3398 | -104.1147
27.3436 | -104.1158
27.3474 | -104.115
27.3511 | -104.1151
27.3549 | -104.1157
27.3587 | -104.1144
27.3625 | -104.1152
27.3662 | -104.115
27.37 | -104.113
27.3738 | -104.1104
27.3776 | -104.1084
27.3813 | -104.1049
27.3851 | -104.1015
27.3889 | -104.0983
27.3927 | -104.0944
27.3964 | -104.0888
27.4002 | -104.0846
27.404 | -104.0822
27.4078 | -104.0792
27.4115 | -104.0756
27.4153 | -104.0702
27.4191 | -104.0637
27.4229 | -104.0592
27.4266 | -104.056
27.4304 | -104.0528
27.4342 | -104.0493
27.438 | -104.0456
27.4417 | -104.0406
27.4455 | -104.0356
27.4493 | -104.0315
27.4531 | -104.0261
27.4568 | -104.0209
27.4606 | -104.0171
27.4644 | -104.0125
27.4682 | -104.0082
27.4719 | -104.0035
27.4757 | -103.999
27.4795 | -103.9961
27.4833 | -103.9948
27.487 | -103.9931
27.4908 | -103.9919
27.4946 | -103.9897
27.4983 | -103.9874
27.5021 | -103.9858
27.5059 | -103.9854
27.5097 | -103.9859
27.5134 | -103.9837
27.5172 | -103.9827
27.521 | -103.982
27.5248 | -103.9801
27.5285 | -103.9796
27.5323 | -103.9787
27.5361 | -103.9788
27.5399 | -103.9812
27.5436 | -103.9817
27.5474 | -103.9815
27.5512 | -103.9839
27.555 | -103.9847
27.5587 | -103.9883
27.5625 | -103.9925
27.5663 | -103.9961
27.5701 | -103.9993
27.5738 | -104.0021
27.5776 | -104.0043
27.5814 | -104.0075
27.5852 | -104.0121
27.5889 | -104.0164
27.5927 | -104.0207
27.5965 | -104.0243
27.6003 | -104.0278
27.604 | -104.0333
27.6078 | -104.0385
27.6116 | -104.0426
27.6154 | -104.0474
27.6191 | -104.0522
27.6229 | -104.0566
27.6267 | -104.0616
27.6305 | -104.0649
27.6342 | -104.0671
27.638 | -104.0688
27.6418 | -104.0718
27.6456 | -104.0773
27.6493 | -104.0837
27.6531 | -104.0888
27.6569 | -104.0936
27.6607 | -104.0973
27.6644 | -104.1014
27.6682 | -104.1075
27.672 | -104.1131
27.6758 | -104.1163
27.6795 | -104.1218
27.6833 | -104.1276
27.6871 | -104.1319
27.6909 | -104.1362
27.6946 | -104.1394
27.6984 | -104.1439
27.7022 | -104.1505
27.706 | -104.156
27.7097 | -104.1598
27.7135 | -104.1632
27.7173 | -104.169
27.7211 | -104.1734
27.7248 | -104.1776
27.7286 | -104.1816
27.7324 | -104.1849
27.7362 | -104.189
27.7399 | -104.193
27.7437 | -104.198
27.7475 | -104.2033
27.7512 | -104.206
27.755 | -104.2096
27.7588 | -104.2138
27.7626 | -104.2173
27.7663 | -104.218
27.7701 | -104.2186
27.7739 | -104.2227
27.7777 | -104.2257
27.7814 | -104.2258
27.7852 | -104.2262
27.789 | -104.2258
27.7928 | -104.2229
27.7965 | -104.2205
27.8003 | -104.2188
27.8041 | -104.218
27.8079 | -104.2181
27.8116 | -104.2181
27.8154 | -104.2148
27.8192 | -104.2104
27.823 | -104.2061
27.8267 | -104.203
27.8305 | -104.1976
27.8343 | -104.1909
27.8381 | -104.1855
27.8418 | -104.1805
27.8456 | -104.1749
27.8494 | -104.1684
27.8532 | -104.1595
27.8569 | -104.1516
27.8607 | -104.1433
27.8645 | -104.1348
27.8683 | -104.1251
27.872 | -104.1157
27.8758 | -104.104
27.8796 | -104.0921
27.8834 | -104.0802
27.8871 | -104.0687
27.8909 | -104.058
27.8947 | -104.0473
27.8985 | -104.0356
27.9022 | -104.0231
27.906 | -104.011
27.9098 | -103.9999
27.9136 | -103.9893
27.9173 | -103.9784
27.9211 | -103.9672
27.9249 | -103.9548
27.9287 | -103.9372
27.9324 | -103.9185
27.9362 | -103.9039
27.94 | -103.8919
27.9438 | -103.8796
27.9475 | -103.8666
27.9513 | -103.8532
27.9551 | -103.8402
27.9589 | -103.8269
27.9626 | -103.8164
27.9664 | -103.8068
27.9702 | -103.7969
27.974 | -103.7891
27.9777 | -103.7801
27.9815 | -103.7701
27.9853 | -103.7619
27.9891 | -103.7538
27.9928 | -103.7453
27.9966 | -103.7371
28.0004 | -103.7304