## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 222 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 221
Hood | 02-92B/P | 159
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1015.1 | hPa |  | PRC PRO 01
Humidity | 41.1 | %/Hr
Internal temperature of calibration bench | 23.91 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1019 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 23.7 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.94 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.01 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3144 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 04/05/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 14/05/2018
Engineer | Anthony HUE | Approval |  | 18/05/2018
CEO | Robert GLEMOT | Approval |  | 18/05/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.389 | 57.3564
0.012209 | -35.9409 | 38.2978
0.018313 | -34.8997 | 28.3234
0.024417 | -34.4548 | 21.9228
0.030521 | -34.3517 | 18.1401
0.036626 | -34.2131 | 15.045
0.04273 | -34.1326 | 12.73
0.048834 | -34.0815 | 11.1423
0.054938 | -34.0139 | 9.9384
0.061043 | -33.9423 | 8.809
0.067147 | -33.9331 | 7.9339
0.073251 | -33.9348 | 7.179
0.079355 | -33.9532 | 6.6219
0.08546 | -33.9723 | 6.1654
0.091564 | -33.9704 | 5.7906
0.097668 | -33.9372 | 5.5334
0.10377 | -33.8904 | 5.1827
0.10988 | -33.8605 | 4.8167
0.11598 | -33.8526 | 4.4543
0.12209 | -33.8706 | 4.1754
0.12819 | -33.9047 | 3.9457
0.13429 | -33.9378 | 3.7105
0.1404 | -33.9558 | 3.44
0.1465 | -33.9723 | 3.184
0.15261 | -33.9801 | 2.945
0.15871 | -33.9773 | 2.7476
0.16482 | -33.9635 | 2.586
0.17092 | -33.9395 | 2.5196
0.17702 | -33.9212 | 2.4191
0.18313 | -33.9045 | 2.3347
0.18923 | -33.8928 | 2.229
0.19534 | -33.8869 | 2.127
0.20144 | -33.8869 | 2.0157
0.20754 | -33.8905 | 1.9425
0.21365 | -33.8936 | 1.8684
0.21975 | -33.8985 | 1.8038
0.22586 | -33.9069 | 1.7386
0.23196 | -33.9158 | 1.6656
0.23807 | -33.9246 | 1.5873
0.24417 | -33.9269 | 1.4927
0.25027 | -33.9305 | 1.4073
0.25638 | -33.9307 | 1.3141
0.26248 | -33.9291 | 1.2261
0.26859 | -33.927 | 1.142
0.27469 | -33.9267 | 1.0553
0.2808 | -33.9331 | 0.99222
0.2869 | -33.9381 | 0.92471
0.293 | -33.944 | 0.86871
0.29911 | -33.9509 | 0.8315
0.30521 | -33.9574 | 0.79639
0.31132 | -33.9618 | 0.78318
0.31742 | -33.9666 | 0.76148
0.32353 | -33.9704 | 0.74348
0.32963 | -33.9737 | 0.72825
0.33573 | -33.9759 | 0.7143
0.34184 | -33.9788 | 0.69705
0.34794 | -33.9822 | 0.66071
0.35405 | -33.9858 | 0.62968
0.36015 | -33.9897 | 0.59538
0.36626 | -33.9923 | 0.5576
0.37236 | -33.995 | 0.50744
0.37846 | -33.9974 | 0.4504
0.38457 | -33.9967 | 0.39281
0.39067 | -33.9976 | 0.34332
0.39678 | -33.9962 | 0.28808
0.40288 | -33.9942 | 0.23618
0.40899 | -33.9907 | 0.18563
0.41509 | -33.9865 | 0.13467
0.42119 | -33.9818 | 0.088034
0.4273 | -33.9783 | 0.039878
0.4334 | -33.9752 | -0.0091168
0.43951 | -33.9724 | -0.057588
0.44561 | -33.9711 | -0.10307
0.45172 | -33.9728 | -0.14124
0.45782 | -33.9728 | -0.18327
0.46392 | -33.9743 | -0.22189
0.47003 | -33.9761 | -0.2572
0.47613 | -33.9784 | -0.28931
0.48224 | -33.9809 | -0.31979
0.48834 | -33.9836 | -0.34099
0.49445 | -33.9859 | -0.36547
0.50055 | -33.9887 | -0.38782
0.50665 | -33.9908 | -0.40575
0.51276 | -33.9923 | -0.42508
0.51886 | -33.9937 | -0.44182
0.52497 | -33.9931 | -0.44804
0.53107 | -33.9937 | -0.45788
0.53717 | -33.9944 | -0.46864
0.54328 | -33.9944 | -0.47354
0.54938 | -33.9946 | -0.4815
0.55549 | -33.9944 | -0.4904
0.56159 | -33.9944 | -0.5019
0.5677 | -33.9942 | -0.51265
0.5738 | -33.9939 | -0.52609
0.5799 | -33.9941 | -0.5384
0.58601 | -33.9948 | -0.55544
0.59211 | -33.996 | -0.57884
0.59822 | -33.9968 | -0.59616
0.60432 | -33.998 | -0.61826
0.61043 | -33.9991 | -0.64022
0.61653 | -34.0005 | -0.66234
0.62263 | -34.0018 | -0.68608
0.62874 | -34.0033 | -0.70685
0.63484 | -34.0047 | -0.72862
0.64095 | -34.0061 | -0.75019
0.64705 | -34.0075 | -0.77075
0.65316 | -34.0087 | -0.79068
0.65926 | -34.01 | -0.81347
0.66536 | -34.0097 | -0.83196
0.67147 | -34.0102 | -0.84862
0.67757 | -34.011 | -0.86692
0.68368 | -34.011 | -0.88697
0.68978 | -34.0114 | -0.90512
0.69589 | -34.0112 | -0.92343
0.70199 | -34.0107 | -0.94518
0.70809 | -34.0103 | -0.96536
0.7142 | -34.0101 | -0.98453
0.7203 | -34.0095 | -1.0041
0.72641 | -34.0089 | -1.0248
0.73251 | -34.0089 | -1.0455
0.73862 | -34.009 | -1.0697
0.74472 | -34.0095 | -1.089
0.75082 | -34.0093 | -1.1099
0.75693 | -34.0095 | -1.1332
0.76303 | -34.0096 | -1.1555
0.76914 | -34.0102 | -1.1825
0.77524 | -34.0105 | -1.2075
0.78135 | -34.011 | -1.2309
0.78745 | -34.0118 | -1.2557
0.79355 | -34.0124 | -1.2804
0.79966 | -34.0134 | -1.3062
0.80576 | -34.0146 | -1.3368
0.81187 | -34.015 | -1.3663
0.81797 | -34.0157 | -1.397
0.82408 | -34.0164 | -1.4242
0.83018 | -34.017 | -1.4554
0.83628 | -34.0179 | -1.4861
0.84239 | -34.0179 | -1.5142
0.84849 | -34.0179 | -1.5422
0.8546 | -34.0184 | -1.5699
0.8607 | -34.0185 | -1.5991
0.86681 | -34.0187 | -1.6275
0.87291 | -34.0184 | -1.6532
0.87901 | -34.0179 | -1.6754
0.88512 | -34.0179 | -1.6984
0.89122 | -34.0178 | -1.7217
0.89733 | -34.0177 | -1.741
0.90343 | -34.0175 | -1.7621
0.90953 | -34.0178 | -1.7769
0.91564 | -34.0174 | -1.7928
0.92174 | -34.0171 | -1.8069
0.92785 | -34.0167 | -1.8202
0.93395 | -34.0169 | -1.8318
0.94006 | -34.017 | -1.8454
0.94616 | -34.017 | -1.8498
0.95226 | -34.0172 | -1.8612
0.95837 | -34.0175 | -1.8727
0.96447 | -34.018 | -1.8827
0.97058 | -34.0182 | -1.8924
0.97668 | -34.0181 | -1.9001
0.98279 | -34.0183 | -1.9034
0.98889 | -34.0187 | -1.9081
0.99499 | -34.0189 | -1.9164
1.0011 | -34.0193 | -1.9211
1.0072 | -34.0193 | -1.9246
1.0133 | -34.0195 | -1.9282
1.0194 | -34.0196 | -1.9335
1.0255 | -34.0194 | -1.94
1.0316 | -34.019 | -1.9439
1.0377 | -34.0186 | -1.9495
1.0438 | -34.018 | -1.9537
1.0499 | -34.0166 | -1.9598
1.056 | -34.0156 | -1.9643
1.0621 | -34.0146 | -1.9697
1.0682 | -34.0143 | -1.9744
1.0743 | -34.0133 | -1.982
1.0805 | -34.0128 | -1.9885
1.0866 | -34.0113 | -1.9991
1.0927 | -34.01 | -2.0063
1.0988 | -34.0094 | -2.0167
1.1049 | -34.008 | -2.024
1.111 | -34.0066 | -2.0349
1.1171 | -34.0053 | -2.044
1.1232 | -34.0041 | -2.0581
1.1293 | -34.0031 | -2.0699
1.1354 | -34.0022 | -2.0822
1.1415 | -34.0011 | -2.096
1.1476 | -33.9998 | -2.1099
1.1537 | -33.9986 | -2.1231
1.1598 | -33.9985 | -2.1414
1.1659 | -33.9981 | -2.1534
1.172 | -33.9974 | -2.1666
1.1781 | -33.9968 | -2.1806
1.1842 | -33.9963 | -2.1962
1.1903 | -33.9971 | -2.2115
1.1964 | -33.9971 | -2.226
1.2025 | -33.9971 | -2.2413
1.2086 | -33.9977 | -2.2571
1.2147 | -33.9981 | -2.2722
1.2209 | -33.9989 | -2.2858
1.227 | -34.0008 | -2.3011
1.2331 | -34.0019 | -2.3154
1.2392 | -34.003 | -2.3314
1.2453 | -34.0043 | -2.3436
1.2514 | -34.0057 | -2.3559
1.2575 | -34.0074 | -2.3701
1.2636 | -34.0098 | -2.3818
1.2697 | -34.0116 | -2.3937
1.2758 | -34.0133 | -2.408
1.2819 | -34.0151 | -2.4217
1.288 | -34.0169 | -2.4356
1.2941 | -34.0191 | -2.4465
1.3002 | -34.0212 | -2.455
1.3063 | -34.023 | -2.4652
1.3124 | -34.0249 | -2.4753
1.3185 | -34.027 | -2.4862
1.3246 | -34.0289 | -2.4972
1.3307 | -34.0311 | -2.5068
1.3368 | -34.0331 | -2.5145
1.3429 | -34.0348 | -2.5234
1.349 | -34.0364 | -2.532
1.3551 | -34.0382 | -2.543
1.3613 | -34.0398 | -2.551
1.3674 | -34.0412 | -2.5595
1.3735 | -34.0426 | -2.5676
1.3796 | -34.0439 | -2.5751
1.3857 | -34.0451 | -2.5821
1.3918 | -34.0462 | -2.5919
1.3979 | -34.0476 | -2.6004
1.404 | -34.0483 | -2.6104
1.4101 | -34.0494 | -2.6193
1.4162 | -34.0503 | -2.6279
1.4223 | -34.0512 | -2.6382
1.4284 | -34.052 | -2.6465
1.4345 | -34.0527 | -2.6567
1.4406 | -34.0533 | -2.6677
1.4467 | -34.0541 | -2.6771
1.4528 | -34.0548 | -2.6862
1.4589 | -34.0557 | -2.6967
1.465 | -34.056 | -2.7044
1.4711 | -34.057 | -2.7133
1.4772 | -34.0576 | -2.7253
1.4833 | -34.0583 | -2.7375
1.4894 | -34.0591 | -2.7488
1.4955 | -34.0599 | -2.7603
1.5016 | -34.0606 | -2.7723
1.5078 | -34.0611 | -2.7867
1.5139 | -34.0619 | -2.7993
1.52 | -34.0624 | -2.8124
1.5261 | -34.0629 | -2.824
1.5322 | -34.0632 | -2.8365
1.5383 | -34.0636 | -2.8486
1.5444 | -34.0643 | -2.8619
1.5505 | -34.0647 | -2.875
1.5566 | -34.0653 | -2.8878
1.5627 | -34.066 | -2.901
1.5688 | -34.0668 | -2.9141
1.5749 | -34.0671 | -2.9263
1.581 | -34.0676 | -2.943
1.5871 | -34.0682 | -2.9586
1.5932 | -34.0687 | -2.9714
1.5993 | -34.0691 | -2.9881
1.6054 | -34.0695 | -3.0031
1.6115 | -34.0699 | -3.0199
1.6176 | -34.0704 | -3.0364
1.6237 | -34.0707 | -3.0507
1.6298 | -34.071 | -3.0661
1.6359 | -34.0715 | -3.0806
1.642 | -34.072 | -3.0952
1.6482 | -34.0725 | -3.1122
1.6543 | -34.073 | -3.129
1.6604 | -34.0735 | -3.1437
1.6665 | -34.0737 | -3.1587
1.6726 | -34.074 | -3.1749
1.6787 | -34.0746 | -3.192
1.6848 | -34.0745 | -3.2085
1.6909 | -34.075 | -3.2248
1.697 | -34.0753 | -3.2411
1.7031 | -34.0754 | -3.2569
1.7092 | -34.0755 | -3.2718
1.7153 | -34.0757 | -3.2881
1.7214 | -34.0754 | -3.3036
1.7275 | -34.0753 | -3.3195
1.7336 | -34.0751 | -3.3357
1.7397 | -34.0752 | -3.3511
1.7458 | -34.0754 | -3.3678
1.7519 | -34.0752 | -3.384
1.758 | -34.0752 | -3.399
1.7641 | -34.0753 | -3.4136
1.7702 | -34.0752 | -3.4294
1.7763 | -34.0753 | -3.4431
1.7824 | -34.0752 | -3.4594
1.7885 | -34.075 | -3.4735
1.7947 | -34.0752 | -3.4885
1.8008 | -34.0752 | -3.5014
1.8069 | -34.0751 | -3.5151
1.813 | -34.0752 | -3.5294
1.8191 | -34.075 | -3.5437
1.8252 | -34.0748 | -3.5576
1.8313 | -34.0749 | -3.5721
1.8374 | -34.075 | -3.5856
1.8435 | -34.0751 | -3.5978
1.8496 | -34.0748 | -3.6123
1.8557 | -34.0748 | -3.6271
1.8618 | -34.0748 | -3.6389
1.8679 | -34.0749 | -3.652
1.874 | -34.0749 | -3.6652
1.8801 | -34.0749 | -3.6782
1.8862 | -34.0751 | -3.691
1.8923 | -34.075 | -3.7032
1.8984 | -34.0748 | -3.7139
1.9045 | -34.0747 | -3.7249
1.9106 | -34.0745 | -3.7368
1.9167 | -34.0743 | -3.7481
1.9228 | -34.0743 | -3.7589
1.9289 | -34.074 | -3.7678
1.9351 | -34.0741 | -3.7769
1.9412 | -34.0742 | -3.7872
1.9473 | -34.0744 | -3.7992
1.9534 | -34.0743 | -3.8104
1.9595 | -34.0741 | -3.8219
1.9656 | -34.0737 | -3.8326
1.9717 | -34.0735 | -3.8431
1.9778 | -34.0736 | -3.8547
1.9839 | -34.0736 | -3.8661
1.99 | -34.0737 | -3.876
1.9961 | -34.0735 | -3.8864
2.0022 | -34.0735 | -3.8964
2.0083 | -34.0732 | -3.9063
2.0144 | -34.0732 | -3.9162
2.0205 | -34.0733 | -3.9281
2.0266 | -34.0735 | -3.9397
2.0327 | -34.0736 | -3.9508
2.0388 | -34.0738 | -3.9647
2.0449 | -34.0739 | -3.9751
2.051 | -34.074 | -3.9851
2.0571 | -34.0739 | -3.9965
2.0632 | -34.0739 | -4.0071
2.0693 | -34.0739 | -4.0182
2.0754 | -34.0742 | -4.0301
2.0816 | -34.0744 | -4.0421
2.0877 | -34.0743 | -4.0532
2.0938 | -34.0746 | -4.0656
2.0999 | -34.0746 | -4.0772
2.106 | -34.0749 | -4.0892
2.1121 | -34.075 | -4.1015
2.1182 | -34.0752 | -4.1116
2.1243 | -34.0754 | -4.1222
2.1304 | -34.0755 | -4.1325
2.1365 | -34.0759 | -4.1441
2.1426 | -34.0761 | -4.1574
2.1487 | -34.0766 | -4.1692
2.1548 | -34.0771 | -4.1822
2.1609 | -34.0777 | -4.1944
2.167 | -34.0782 | -4.207
2.1731 | -34.0785 | -4.219
2.1792 | -34.0793 | -4.2303
2.1853 | -34.0798 | -4.2432
2.1914 | -34.0804 | -4.2555
2.1975 | -34.081 | -4.2676
2.2036 | -34.0815 | -4.2809
2.2097 | -34.0819 | -4.293
2.2158 | -34.083 | -4.3033
2.222 | -34.0835 | -4.3136
2.2281 | -34.0839 | -4.3244
2.2342 | -34.0845 | -4.3376
2.2403 | -34.085 | -4.3477
2.2464 | -34.0858 | -4.3589
2.2525 | -34.0863 | -4.3698
2.2586 | -34.0873 | -4.3797
2.2647 | -34.0881 | -4.3891
2.2708 | -34.0888 | -4.4009
2.2769 | -34.0897 | -4.4125
2.283 | -34.0902 | -4.422
2.2891 | -34.0908 | -4.4324
2.2952 | -34.0915 | -4.4423
2.3013 | -34.0922 | -4.4526
2.3074 | -34.0929 | -4.4623
2.3135 | -34.0934 | -4.4737
2.3196 | -34.0941 | -4.4824
2.3257 | -34.0945 | -4.493
2.3318 | -34.0949 | -4.5021
2.3379 | -34.0955 | -4.5137
2.344 | -34.0959 | -4.5246
2.3501 | -34.0964 | -4.5354
2.3562 | -34.0967 | -4.5438
2.3623 | -34.0971 | -4.5526
2.3685 | -34.0974 | -4.5617
2.3746 | -34.0977 | -4.5723
2.3807 | -34.0982 | -4.5813
2.3868 | -34.0985 | -4.5904
2.3929 | -34.0989 | -4.5998
2.399 | -34.0992 | -4.6079
2.4051 | -34.0998 | -4.6175
2.4112 | -34.1001 | -4.6266
2.4173 | -34.1005 | -4.6356
2.4234 | -34.1006 | -4.6442
2.4295 | -34.1009 | -4.6528
2.4356 | -34.1011 | -4.6622
2.4417 | -34.1012 | -4.6723
2.4478 | -34.1013 | -4.682
2.4539 | -34.1017 | -4.6899
2.46 | -34.1014 | -4.6997
2.4661 | -34.1015 | -4.709
2.4722 | -34.1018 | -4.7178
2.4783 | -34.1019 | -4.7263
2.4844 | -34.102 | -4.7349
2.4905 | -34.1019 | -4.7442
2.4966 | -34.1017 | -4.7531
2.5027 | -34.1017 | -4.7612
2.5089 | -34.1017 | -4.7698
2.515 | -34.1015 | -4.7791
2.5211 | -34.1015 | -4.7875
2.5272 | -34.1012 | -4.7966
2.5333 | -34.1011 | -4.8062
2.5394 | -34.1011 | -4.8153
2.5455 | -34.101 | -4.8247
2.5516 | -34.101 | -4.8335
2.5577 | -34.1009 | -4.8435
2.5638 | -34.1008 | -4.8539
2.5699 | -34.1006 | -4.8642
2.576 | -34.1004 | -4.8746
2.5821 | -34.1002 | -4.8846
2.5882 | -34.1001 | -4.8942
2.5943 | -34.1 | -4.9041
2.6004 | -34.0998 | -4.9141
2.6065 | -34.0997 | -4.9227
2.6126 | -34.0996 | -4.9319
2.6187 | -34.0996 | -4.9407
2.6248 | -34.0995 | -4.9498
2.6309 | -34.0994 | -4.9611
2.637 | -34.0994 | -4.9723
2.6431 | -34.0994 | -4.9819
2.6492 | -34.0993 | -4.992
2.6554 | -34.0994 | -5.0023
2.6615 | -34.0992 | -5.0122
2.6676 | -34.0992 | -5.0227
2.6737 | -34.0994 | -5.0335
2.6798 | -34.0994 | -5.0435
2.6859 | -34.0994 | -5.0541
2.692 | -34.0994 | -5.0639
2.6981 | -34.0994 | -5.0751
2.7042 | -34.0995 | -5.0873
2.7103 | -34.0995 | -5.0974
2.7164 | -34.0997 | -5.1082
2.7225 | -34.0999 | -5.1196
2.7286 | -34.1 | -5.1309
2.7347 | -34.1 | -5.1409
2.7408 | -34.1004 | -5.1502
2.7469 | -34.1003 | -5.1617
2.753 | -34.1007 | -5.1725
2.7591 | -34.1009 | -5.1845
2.7652 | -34.1009 | -5.1961
2.7713 | -34.1009 | -5.206
2.7774 | -34.1014 | -5.2175
2.7835 | -34.1015 | -5.2283
2.7896 | -34.1017 | -5.2377
2.7958 | -34.1021 | -5.248
2.8019 | -34.1027 | -5.2579
2.808 | -34.103 | -5.2699
2.8141 | -34.1035 | -5.283
2.8202 | -34.1039 | -5.2952
2.8263 | -34.1044 | -5.3063
2.8324 | -34.1046 | -5.3188
2.8385 | -34.1051 | -5.331
2.8446 | -34.1057 | -5.3428
2.8507 | -34.1061 | -5.355
2.8568 | -34.1066 | -5.3663
2.8629 | -34.107 | -5.3777
2.869 | -34.1075 | -5.3891
2.8751 | -34.108 | -5.4012
2.8812 | -34.1087 | -5.4146
2.8873 | -34.109 | -5.426
2.8934 | -34.1094 | -5.438
2.8995 | -34.1098 | -5.4501
2.9056 | -34.1102 | -5.4609
2.9117 | -34.1107 | -5.4726
2.9178 | -34.1114 | -5.4854
2.9239 | -34.1119 | -5.4988
2.93 | -34.1124 | -5.5124
2.9361 | -34.1129 | -5.5239
2.9423 | -34.1136 | -5.536
2.9484 | -34.1141 | -5.5497
2.9545 | -34.1148 | -5.563
2.9606 | -34.1154 | -5.5765
2.9667 | -34.1159 | -5.5884
2.9728 | -34.1164 | -5.6008
2.9789 | -34.1169 | -5.6136
2.985 | -34.1173 | -5.6275
2.9911 | -34.1178 | -5.6402
2.9972 | -34.1184 | -5.653
3.0033 | -34.119 | -5.6662
3.0094 | -34.1196 | -5.6793
3.0155 | -34.1202 | -5.692
3.0216 | -34.1208 | -5.7056
3.0277 | -34.1212 | -5.7194
3.0338 | -34.1219 | -5.7331
3.0399 | -34.1225 | -5.7469
3.046 | -34.1229 | -5.76
3.0521 | -34.1235 | -5.7726
3.0582 | -34.1238 | -5.7866
3.0643 | -34.1244 | -5.7997
3.0704 | -34.1247 | -5.8132
3.0765 | -34.1251 | -5.8255
3.0827 | -34.1254 | -5.8387
3.0888 | -34.1258 | -5.8517
3.0949 | -34.1262 | -5.8648
3.101 | -34.1267 | -5.879
3.1071 | -34.127 | -5.8932
3.1132 | -34.1275 | -5.9074
3.1193 | -34.128 | -5.9205
3.1254 | -34.1284 | -5.9328
3.1315 | -34.1288 | -5.9472
3.1376 | -34.1291 | -5.9602
3.1437 | -34.1295 | -5.9738
3.1498 | -34.1299 | -5.9881
3.1559 | -34.1302 | -6.0014
3.162 | -34.1303 | -6.0156
3.1681 | -34.1304 | -6.0299
3.1742 | -34.1308 | -6.0437
3.1803 | -34.1312 | -6.057
3.1864 | -34.1316 | -6.0693
3.1925 | -34.132 | -6.0831
3.1986 | -34.1321 | -6.0959
3.2047 | -34.1322 | -6.11
3.2108 | -34.1325 | -6.1218
3.2169 | -34.133 | -6.1357
3.223 | -34.133 | -6.1491
3.2292 | -34.133 | -6.1631
3.2353 | -34.1329 | -6.1727
3.2414 | -34.1331 | -6.1858
3.2475 | -34.1331 | -6.1969
3.2536 | -34.1334 | -6.2107
3.2597 | -34.1337 | -6.222
3.2658 | -34.1339 | -6.2328
3.2719 | -34.134 | -6.244
3.278 | -34.1341 | -6.2562
3.2841 | -34.1343 | -6.2682
3.2902 | -34.1344 | -6.2808
3.2963 | -34.1346 | -6.2925
3.3024 | -34.1345 | -6.3028
3.3085 | -34.1348 | -6.3154
3.3146 | -34.1347 | -6.3264
3.3207 | -34.135 | -6.3386
3.3268 | -34.1351 | -6.3509
3.3329 | -34.1353 | -6.3607
3.339 | -34.1354 | -6.3704
3.3451 | -34.1355 | -6.3819
3.3512 | -34.1356 | -6.3914
3.3573 | -34.1358 | -6.4028
3.3634 | -34.1359 | -6.4144
3.3696 | -34.1359 | -6.4263
3.3757 | -34.1359 | -6.4362
3.3818 | -34.1363 | -6.448
3.3879 | -34.1362 | -6.4584
3.394 | -34.1362 | -6.4689
3.4001 | -34.1363 | -6.4798
3.4062 | -34.1363 | -6.49
3.4123 | -34.1364 | -6.4999
3.4184 | -34.1365 | -6.5109
3.4245 | -34.1366 | -6.5215
3.4306 | -34.1366 | -6.5313
3.4367 | -34.1368 | -6.5411
3.4428 | -34.1371 | -6.5503
3.4489 | -34.1371 | -6.5615
3.455 | -34.137 | -6.572
3.4611 | -34.1371 | -6.5831
3.4672 | -34.1371 | -6.5945
3.4733 | -34.1372 | -6.6047
3.4794 | -34.1372 | -6.6152
3.4855 | -34.1375 | -6.6249
3.4916 | -34.1377 | -6.6346
3.4977 | -34.1376 | -6.6456
3.5038 | -34.1378 | -6.6556
3.5099 | -34.1379 | -6.6652
3.5161 | -34.1381 | -6.6748
3.5222 | -34.1382 | -6.6849
3.5283 | -34.1382 | -6.695
3.5344 | -34.1383 | -6.7058
3.5405 | -34.1384 | -6.7157
3.5466 | -34.1386 | -6.7255
3.5527 | -34.1387 | -6.7362
3.5588 | -34.1388 | -6.746
3.5649 | -34.1389 | -6.7562
3.571 | -34.139 | -6.7665
3.5771 | -34.1392 | -6.7773
3.5832 | -34.1393 | -6.787
3.5893 | -34.1396 | -6.7968
3.5954 | -34.1397 | -6.8074
3.6015 | -34.1399 | -6.8178
3.6076 | -34.1402 | -6.8274
3.6137 | -34.1403 | -6.8374
3.6198 | -34.1407 | -6.8464
3.6259 | -34.1408 | -6.8572
3.632 | -34.141 | -6.8679
3.6381 | -34.1413 | -6.8778
3.6442 | -34.1415 | -6.8882
3.6503 | -34.1418 | -6.8973
3.6565 | -34.1421 | -6.9091
3.6626 | -34.1422 | -6.9199
3.6687 | -34.1425 | -6.9305
3.6748 | -34.1429 | -6.9405
3.6809 | -34.1433 | -6.951
3.687 | -34.1435 | -6.9622
3.6931 | -34.1439 | -6.9723
3.6992 | -34.1442 | -6.9829
3.7053 | -34.1444 | -6.9931
3.7114 | -34.1446 | -7.0044
3.7175 | -34.145 | -7.0143
3.7236 | -34.1453 | -7.0231
3.7297 | -34.1454 | -7.0339
3.7358 | -34.1457 | -7.0457
3.7419 | -34.1459 | -7.0573
3.748 | -34.1461 | -7.0681
3.7541 | -34.1463 | -7.0791
3.7602 | -34.1467 | -7.09
3.7663 | -34.1469 | -7.1002
3.7724 | -34.1473 | -7.1113
3.7785 | -34.1477 | -7.1217
3.7846 | -34.1479 | -7.1324
3.7907 | -34.1483 | -7.1434
3.7969 | -34.1486 | -7.1549
3.803 | -34.149 | -7.165
3.8091 | -34.1493 | -7.1754
3.8152 | -34.1496 | -7.1866
3.8213 | -34.1498 | -7.1982
3.8274 | -34.1503 | -7.2088
3.8335 | -34.1507 | -7.2204
3.8396 | -34.1512 | -7.232
3.8457 | -34.1515 | -7.2422
3.8518 | -34.1521 | -7.2531
3.8579 | -34.1525 | -7.2638
3.864 | -34.1529 | -7.2745
3.8701 | -34.1533 | -7.2855
3.8762 | -34.1536 | -7.2967
3.8823 | -34.1539 | -7.3074
3.8884 | -34.1544 | -7.3178
3.8945 | -34.1549 | -7.3292
3.9006 | -34.1553 | -7.3407
3.9067 | -34.1557 | -7.3526
3.9128 | -34.1561 | -7.364
3.9189 | -34.1565 | -7.3746
3.925 | -34.1568 | -7.3855
3.9311 | -34.1573 | -7.3956
3.9372 | -34.1577 | -7.4072
3.9434 | -34.1581 | -7.4177
3.9495 | -34.1585 | -7.4293
3.9556 | -34.159 | -7.4406
3.9617 | -34.1593 | -7.4515
3.9678 | -34.1597 | -7.4623
3.9739 | -34.1602 | -7.4735
3.98 | -34.1606 | -7.4854
3.9861 | -34.1609 | -7.4959
3.9922 | -34.1612 | -7.5076
3.9983 | -34.1616 | -7.5177
4.0044 | -34.1619 | -7.5275
4.0105 | -34.1623 | -7.5384
4.0166 | -34.1629 | -7.5495
4.0227 | -34.1633 | -7.5599
4.0288 | -34.1635 | -7.5713
4.0349 | -34.1639 | -7.5812
4.041 | -34.1643 | -7.5922
4.0471 | -34.1645 | -7.6024
4.0532 | -34.165 | -7.6137
4.0593 | -34.1653 | -7.6253
4.0654 | -34.1657 | -7.6361
4.0715 | -34.1661 | -7.6475
4.0776 | -34.1664 | -7.6579
4.0838 | -34.1668 | -7.6693
4.0899 | -34.1672 | -7.6806
4.096 | -34.1675 | -7.6911
4.1021 | -34.1679 | -7.7019
4.1082 | -34.1683 | -7.7122
4.1143 | -34.1687 | -7.7222
4.1204 | -34.1691 | -7.733
4.1265 | -34.1695 | -7.7435
4.1326 | -34.1699 | -7.7538
4.1387 | -34.1704 | -7.7651
4.1448 | -34.1708 | -7.7761
4.1509 | -34.171 | -7.7868
4.157 | -34.1713 | -7.7981
4.1631 | -34.1716 | -7.8093
4.1692 | -34.172 | -7.8201
4.1753 | -34.1724 | -7.831
4.1814 | -34.1727 | -7.8408
4.1875 | -34.1731 | -7.8523
4.1936 | -34.1734 | -7.8637
4.1997 | -34.1737 | -7.8738
4.2058 | -34.174 | -7.8854
4.2119 | -34.1743 | -7.8952
4.218 | -34.1746 | -7.9054
4.2241 | -34.175 | -7.915
4.2303 | -34.1755 | -7.9263
4.2364 | -34.1758 | -7.9378
4.2425 | -34.1761 | -7.9483
4.2486 | -34.1764 | -7.9595
4.2547 | -34.1766 | -7.9688
4.2608 | -34.1769 | -7.9788
4.2669 | -34.1772 | -7.9892
4.273 | -34.1775 | -8
4.2791 | -34.1778 | -8.0108
4.2852 | -34.178 | -8.0211
4.2913 | -34.1782 | -8.0304
4.2974 | -34.1784 | -8.0414
4.3035 | -34.1787 | -8.0522
4.3096 | -34.1789 | -8.0627
4.3157 | -34.1792 | -8.0727
4.3218 | -34.1794 | -8.0827
4.3279 | -34.1797 | -8.0932
4.334 | -34.1799 | -8.1031
4.3401 | -34.1802 | -8.113
4.3462 | -34.1805 | -8.1239
4.3523 | -34.1809 | -8.1337
4.3584 | -34.181 | -8.1422
4.3645 | -34.1812 | -8.1519
4.3707 | -34.1815 | -8.1612
4.3768 | -34.1817 | -8.1717
4.3829 | -34.1819 | -8.182
4.389 | -34.1822 | -8.1918
4.3951 | -34.1825 | -8.2013
4.4012 | -34.1827 | -8.2112
4.4073 | -34.183 | -8.2214
4.4134 | -34.1832 | -8.2316
4.4195 | -34.1835 | -8.2412
4.4256 | -34.1838 | -8.2509
4.4317 | -34.1841 | -8.2604
4.4378 | -34.1843 | -8.27
4.4439 | -34.1845 | -8.2796
4.45 | -34.1846 | -8.2896
4.4561 | -34.1848 | -8.2995
4.4622 | -34.1849 | -8.3091
4.4683 | -34.1851 | -8.3184
4.4744 | -34.1854 | -8.3286
4.4805 | -34.1857 | -8.3378
4.4866 | -34.1859 | -8.3471
4.4927 | -34.1861 | -8.3574
4.4988 | -34.1863 | -8.3668
4.5049 | -34.1866 | -8.376
4.511 | -34.1869 | -8.3851
4.5172 | -34.1872 | -8.395
4.5233 | -34.1874 | -8.4056
4.5294 | -34.1877 | -8.415
4.5355 | -34.1879 | -8.4245
4.5416 | -34.1882 | -8.4336
4.5477 | -34.1884 | -8.443
4.5538 | -34.1886 | -8.4518
4.5599 | -34.189 | -8.4613
4.566 | -34.1892 | -8.4711
4.5721 | -34.1895 | -8.4804
4.5782 | -34.1898 | -8.4902
4.5843 | -34.19 | -8.4999
4.5904 | -34.1904 | -8.5092
4.5965 | -34.1907 | -8.518
4.6026 | -34.191 | -8.5287
4.6087 | -34.1912 | -8.5378
4.6148 | -34.1915 | -8.5485
4.6209 | -34.1919 | -8.557
4.627 | -34.1922 | -8.5671
4.6331 | -34.1923 | -8.5761
4.6392 | -34.1926 | -8.5854
4.6453 | -34.1929 | -8.5948
4.6514 | -34.1932 | -8.6044
4.6576 | -34.1933 | -8.6149
4.6637 | -34.1935 | -8.6226
4.6698 | -34.1938 | -8.6313
4.6759 | -34.1939 | -8.6411
4.682 | -34.1941 | -8.6503
4.6881 | -34.1943 | -8.6592
4.6942 | -34.1946 | -8.6685
4.7003 | -34.1949 | -8.678
4.7064 | -34.1951 | -8.6873
4.7125 | -34.1955 | -8.6968
4.7186 | -34.1957 | -8.7053
4.7247 | -34.196 | -8.715
4.7308 | -34.1964 | -8.7244
4.7369 | -34.1966 | -8.734
4.743 | -34.1968 | -8.7444
4.7491 | -34.197 | -8.7541
4.7552 | -34.1972 | -8.7634
4.7613 | -34.1974 | -8.773
4.7674 | -34.1977 | -8.7817
4.7735 | -34.1979 | -8.7918
4.7796 | -34.1981 | -8.8017
4.7857 | -34.1982 | -8.8114
4.7918 | -34.1985 | -8.8201
4.7979 | -34.1989 | -8.8292
4.8041 | -34.199 | -8.8386
4.8102 | -34.1994 | -8.8477
4.8163 | -34.1996 | -8.8572
4.8224 | -34.1998 | -8.8659
4.8285 | -34.2001 | -8.875
4.8346 | -34.2004 | -8.8847
4.8407 | -34.2006 | -8.894
4.8468 | -34.2009 | -8.9038
4.8529 | -34.2011 | -8.9137
4.859 | -34.2014 | -8.9232
4.8651 | -34.2017 | -8.9326
4.8712 | -34.2019 | -8.9421
4.8773 | -34.2021 | -8.9517
4.8834 | -34.2024 | -8.9615
4.8895 | -34.2027 | -8.9716
4.8956 | -34.2029 | -8.9817
4.9017 | -34.2031 | -8.9912
4.9078 | -34.2033 | -9.0011
4.9139 | -34.2036 | -9.0098
4.92 | -34.2039 | -9.0189
4.9261 | -34.2043 | -9.0278
4.9322 | -34.2046 | -9.037
4.9383 | -34.205 | -9.0461
4.9445 | -34.2053 | -9.0563
4.9506 | -34.2055 | -9.0652
4.9567 | -34.2058 | -9.0745
4.9628 | -34.2061 | -9.084
4.9689 | -34.2063 | -9.0938
4.975 | -34.2065 | -9.1032
4.9811 | -34.2069 | -9.1131
4.9872 | -34.2071 | -9.123
4.9933 | -34.2075 | -9.1327
4.9994 | -34.2079 | -9.1426
5.0055 | -34.2082 | -9.1523
5.0116 | -34.2086 | -9.1626
5.0177 | -34.209 | -9.1728
5.0238 | -34.2093 | -9.1833
5.0299 | -34.2098 | -9.1944
5.036 | -34.2101 | -9.204
5.0421 | -34.2103 | -9.2138
5.0482 | -34.2106 | -9.2241
5.0543 | -34.211 | -9.2335
5.0604 | -34.2114 | -9.2438
5.0665 | -34.2117 | -9.2535
5.0726 | -34.2121 | -9.264
5.0787 | -34.2124 | -9.2743
5.0848 | -34.2128 | -9.2844
5.091 | -34.2131 | -9.2944
5.0971 | -34.2134 | -9.3045
5.1032 | -34.2138 | -9.315
5.1093 | -34.2141 | -9.325
5.1154 | -34.2144 | -9.335
5.1215 | -34.2147 | -9.3451
5.1276 | -34.215 | -9.355
5.1337 | -34.2155 | -9.3648
5.1398 | -34.2158 | -9.3747
5.1459 | -34.2161 | -9.3852
5.152 | -34.2165 | -9.3963
5.1581 | -34.2167 | -9.407
5.1642 | -34.217 | -9.4171
5.1703 | -34.2174 | -9.4271
5.1764 | -34.2178 | -9.4371
5.1825 | -34.2181 | -9.4475
5.1886 | -34.2184 | -9.4574
5.1947 | -34.2187 | -9.4683
5.2008 | -34.2191 | -9.4782
5.2069 | -34.2195 | -9.4894
5.213 | -34.2199 | -9.4991
5.2191 | -34.2203 | -9.5095
5.2252 | -34.2206 | -9.5198
5.2314 | -34.2208 | -9.5306
5.2375 | -34.2213 | -9.5406
5.2436 | -34.2215 | -9.5509
5.2497 | -34.2219 | -9.5608
5.2558 | -34.2223 | -9.5719
5.2619 | -34.2226 | -9.5828
5.268 | -34.223 | -9.5933
5.2741 | -34.2234 | -9.6046
5.2802 | -34.2237 | -9.6155
5.2863 | -34.2242 | -9.6258
5.2924 | -34.2244 | -9.637
5.2985 | -34.2249 | -9.6484
5.3046 | -34.2254 | -9.6598
5.3107 | -34.2258 | -9.6707
5.3168 | -34.2262 | -9.6812
5.3229 | -34.2266 | -9.6911
5.329 | -34.227 | -9.7011
5.3351 | -34.2274 | -9.7114
5.3412 | -34.2279 | -9.7216
5.3473 | -34.2283 | -9.7326
5.3534 | -34.2287 | -9.7435
5.3595 | -34.229 | -9.7543
5.3656 | -34.2294 | -9.7654
5.3717 | -34.2297 | -9.776
5.3779 | -34.2301 | -9.787
5.384 | -34.2304 | -9.7978
5.3901 | -34.2308 | -9.8085
5.3962 | -34.2311 | -9.8194
5.4023 | -34.2315 | -9.8303
5.4084 | -34.232 | -9.8407
5.4145 | -34.2323 | -9.8519
5.4206 | -34.2326 | -9.8632
5.4267 | -34.233 | -9.8747
5.4328 | -34.2334 | -9.8849
5.4389 | -34.2338 | -9.8952
5.445 | -34.2341 | -9.9066
5.4511 | -34.2345 | -9.9171
5.4572 | -34.2349 | -9.9283
5.4633 | -34.2353 | -9.9388
5.4694 | -34.2356 | -9.9494
5.4755 | -34.236 | -9.9603
5.4816 | -34.2363 | -9.971
5.4877 | -34.2367 | -9.9818
5.4938 | -34.237 | -9.9923
5.4999 | -34.2374 | -10.0031
5.506 | -34.2378 | -10.0139
5.5121 | -34.2382 | -10.025
5.5183 | -34.2385 | -10.0352
5.5244 | -34.2388 | -10.0461
5.5305 | -34.2392 | -10.0571
5.5366 | -34.2395 | -10.0676
5.5427 | -34.2399 | -10.0789
5.5488 | -34.2402 | -10.09
5.5549 | -34.2405 | -10.1008
5.561 | -34.2409 | -10.1121
5.5671 | -34.2413 | -10.1226
5.5732 | -34.2416 | -10.1336
5.5793 | -34.242 | -10.1452
5.5854 | -34.2425 | -10.1562
5.5915 | -34.2427 | -10.1673
5.5976 | -34.243 | -10.1783
5.6037 | -34.2434 | -10.1891
5.6098 | -34.2437 | -10.1997
5.6159 | -34.2441 | -10.2107
5.622 | -34.2445 | -10.2212
5.6281 | -34.2448 | -10.2322
5.6342 | -34.2452 | -10.2431
5.6403 | -34.2455 | -10.2534
5.6464 | -34.2459 | -10.2643
5.6525 | -34.2462 | -10.2749
5.6586 | -34.2467 | -10.2849
5.6648 | -34.247 | -10.2963
5.6709 | -34.2473 | -10.3075
5.677 | -34.2477 | -10.3191
5.6831 | -34.248 | -10.3294
5.6892 | -34.2484 | -10.3402
5.6953 | -34.2487 | -10.351
5.7014 | -34.2491 | -10.3623
5.7075 | -34.2494 | -10.3734
5.7136 | -34.2498 | -10.384
5.7197 | -34.2501 | -10.3947
5.7258 | -34.2504 | -10.4053
5.7319 | -34.2507 | -10.4161
5.738 | -34.2511 | -10.4269
5.7441 | -34.2515 | -10.4377
5.7502 | -34.2519 | -10.4481
5.7563 | -34.2522 | -10.4586
5.7624 | -34.2526 | -10.4701
5.7685 | -34.2528 | -10.4807
5.7746 | -34.2531 | -10.4913
5.7807 | -34.2534 | -10.5018
5.7868 | -34.2538 | -10.5125
5.7929 | -34.2541 | -10.5231
5.799 | -34.2544 | -10.5338
5.8052 | -34.2547 | -10.5446
5.8113 | -34.255 | -10.5554
5.8174 | -34.2553 | -10.5663
5.8235 | -34.2556 | -10.5769
5.8296 | -34.256 | -10.5882
5.8357 | -34.2562 | -10.5989
5.8418 | -34.2566 | -10.6099
5.8479 | -34.2568 | -10.6209
5.854 | -34.2571 | -10.632
5.8601 | -34.2575 | -10.6425
5.8662 | -34.2578 | -10.6536
5.8723 | -34.2581 | -10.6637
5.8784 | -34.2584 | -10.6744
5.8845 | -34.2587 | -10.6852
5.8906 | -34.259 | -10.6959
5.8967 | -34.2594 | -10.7062
5.9028 | -34.2597 | -10.717
5.9089 | -34.26 | -10.7271
5.915 | -34.2604 | -10.7378
5.9211 | -34.2607 | -10.7485
5.9272 | -34.261 | -10.7591
5.9333 | -34.2613 | -10.7691
5.9394 | -34.2617 | -10.7795
5.9455 | -34.262 | -10.7898
5.9517 | -34.2624 | -10.8005
5.9578 | -34.2628 | -10.8114
5.9639 | -34.2631 | -10.8217
5.97 | -34.2635 | -10.8325
5.9761 | -34.2638 | -10.8428
5.9822 | -34.2642 | -10.8535
5.9883 | -34.2645 | -10.8639
5.9944 | -34.2649 | -10.8746
6.0005 | -34.2653 | -10.8855
6.0066 | -34.2656 | -10.8957
6.0127 | -34.2659 | -10.9056
6.0188 | -34.2662 | -10.9165
6.0249 | -34.2666 | -10.9271
6.031 | -34.2669 | -10.9374
6.0371 | -34.2673 | -10.9482
6.0432 | -34.2677 | -10.9588
6.0493 | -34.2681 | -10.9691
6.0554 | -34.2685 | -10.9796
6.0615 | -34.2688 | -10.9907
6.0676 | -34.2691 | -11.0018
6.0737 | -34.2695 | -11.0126
6.0798 | -34.2698 | -11.023
6.0859 | -34.2702 | -11.0335
6.0921 | -34.2705 | -11.0438
6.0982 | -34.2709 | -11.0546
6.1043 | -34.2713 | -11.065
6.1104 | -34.2717 | -11.076
6.1165 | -34.2721 | -11.0858
6.1226 | -34.2725 | -11.0957
6.1287 | -34.2728 | -11.1063
6.1348 | -34.2733 | -11.1169
6.1409 | -34.2736 | -11.1273
6.147 | -34.274 | -11.1376
6.1531 | -34.2743 | -11.1481
6.1592 | -34.2747 | -11.1586
6.1653 | -34.275 | -11.1689
6.1714 | -34.2754 | -11.1794
6.1775 | -34.2758 | -11.1902
6.1836 | -34.2761 | -11.2004
6.1897 | -34.2765 | -11.2109
6.1958 | -34.2768 | -11.2216
6.2019 | -34.2772 | -11.2324
6.208 | -34.2776 | -11.2433
6.2141 | -34.278 | -11.2539
6.2202 | -34.2783 | -11.2649
6.2263 | -34.2787 | -11.2752
6.2325 | -34.2791 | -11.2852
6.2386 | -34.2795 | -11.2954
6.2447 | -34.2799 | -11.3061
6.2508 | -34.2802 | -11.3165
6.2569 | -34.2806 | -11.3273
6.263 | -34.281 | -11.3374
6.2691 | -34.2814 | -11.3477
6.2752 | -34.2818 | -11.3581
6.2813 | -34.2821 | -11.3685
6.2874 | -34.2825 | -11.379
6.2935 | -34.2829 | -11.3892
6.2996 | -34.2833 | -11.3997
6.3057 | -34.2837 | -11.4101
6.3118 | -34.284 | -11.4205
6.3179 | -34.2844 | -11.4309
6.324 | -34.2848 | -11.4416
6.3301 | -34.2852 | -11.4518
6.3362 | -34.2856 | -11.4623
6.3423 | -34.2861 | -11.4728
6.3484 | -34.2865 | -11.4834
6.3545 | -34.2869 | -11.4932
6.3606 | -34.2873 | -11.5034
6.3667 | -34.2877 | -11.5135
6.3728 | -34.2882 | -11.524
6.379 | -34.2885 | -11.5344
6.3851 | -34.2889 | -11.5448
6.3912 | -34.2893 | -11.5549
6.3973 | -34.2898 | -11.5651
6.4034 | -34.2901 | -11.5754
6.4095 | -34.2905 | -11.586
6.4156 | -34.2909 | -11.5961
6.4217 | -34.2912 | -11.6064
6.4278 | -34.2916 | -11.6167
6.4339 | -34.2921 | -11.627
6.44 | -34.2925 | -11.637
6.4461 | -34.2928 | -11.6478
6.4522 | -34.2932 | -11.6576
6.4583 | -34.2936 | -11.668
6.4644 | -34.2939 | -11.6787
6.4705 | -34.2944 | -11.6892
6.4766 | -34.2948 | -11.7
6.4827 | -34.2952 | -11.7108
6.4888 | -34.2955 | -11.721
6.4949 | -34.2959 | -11.7309
6.501 | -34.2963 | -11.7416
6.5071 | -34.2967 | -11.752
6.5132 | -34.2971 | -11.7622
6.5194 | -34.2975 | -11.7725
6.5255 | -34.2979 | -11.7825
6.5316 | -34.2983 | -11.793
6.5377 | -34.2987 | -11.8036
6.5438 | -34.2991 | -11.8138
6.5499 | -34.2995 | -11.824
6.556 | -34.2999 | -11.8341
6.5621 | -34.3002 | -11.8446
6.5682 | -34.3006 | -11.8548
6.5743 | -34.301 | -11.8659
6.5804 | -34.3013 | -11.8762
6.5865 | -34.3016 | -11.8868
6.5926 | -34.3021 | -11.8971
6.5987 | -34.3025 | -11.9071
6.6048 | -34.3029 | -11.9181
6.6109 | -34.3033 | -11.9283
6.617 | -34.3037 | -11.9396
6.6231 | -34.304 | -11.9501
6.6292 | -34.3045 | -11.9604
6.6353 | -34.3048 | -11.9717
6.6414 | -34.3053 | -11.9824
6.6475 | -34.3057 | -11.9929
6.6536 | -34.3061 | -12.0029
6.6597 | -34.3065 | -12.013
6.6659 | -34.3069 | -12.0232
6.672 | -34.3074 | -12.0336
6.6781 | -34.3077 | -12.0438
6.6842 | -34.3081 | -12.0543
6.6903 | -34.3086 | -12.0646
6.6964 | -34.309 | -12.075
6.7025 | -34.3093 | -12.0854
6.7086 | -34.3097 | -12.0955
6.7147 | -34.3101 | -12.106
6.7208 | -34.3105 | -12.1169
6.7269 | -34.3108 | -12.1269
6.733 | -34.3111 | -12.1365
6.7391 | -34.3116 | -12.147
6.7452 | -34.3121 | -12.1568
6.7513 | -34.3125 | -12.1668
6.7574 | -34.3129 | -12.1774
6.7635 | -34.3133 | -12.1872
6.7696 | -34.3138 | -12.1976
6.7757 | -34.3142 | -12.2083
6.7818 | -34.3146 | -12.2186
6.7879 | -34.315 | -12.2291
6.794 | -34.3155 | -12.2392
6.8001 | -34.3158 | -12.2493
6.8063 | -34.3162 | -12.2598
6.8124 | -34.3166 | -12.2703
6.8185 | -34.317 | -12.2805
6.8246 | -34.3174 | -12.2909
6.8307 | -34.3178 | -12.301
6.8368 | -34.3182 | -12.3116
6.8429 | -34.3186 | -12.3219
6.849 | -34.319 | -12.3322
6.8551 | -34.3194 | -12.3425
6.8612 | -34.3198 | -12.3526
6.8673 | -34.3202 | -12.3627
6.8734 | -34.3206 | -12.3738
6.8795 | -34.3211 | -12.3843
6.8856 | -34.3214 | -12.3941
6.8917 | -34.3219 | -12.4038
6.8978 | -34.3222 | -12.414
6.9039 | -34.3226 | -12.4241
6.91 | -34.323 | -12.4345
6.9161 | -34.3234 | -12.4446
6.9222 | -34.3238 | -12.4549
6.9283 | -34.3242 | -12.4648
6.9344 | -34.3245 | -12.4749
6.9405 | -34.3249 | -12.4849
6.9466 | -34.3253 | -12.4954
6.9528 | -34.3257 | -12.5057
6.9589 | -34.3261 | -12.5158
6.965 | -34.3265 | -12.526
6.9711 | -34.327 | -12.5363
6.9772 | -34.3274 | -12.5463
6.9833 | -34.3277 | -12.5567
6.9894 | -34.3281 | -12.5673
6.9955 | -34.3285 | -12.5775
7.0016 | -34.3289 | -12.5881
7.0077 | -34.3293 | -12.5987
7.0138 | -34.3297 | -12.6092
7.0199 | -34.3301 | -12.6192
7.026 | -34.3304 | -12.6294
7.0321 | -34.3307 | -12.6396
7.0382 | -34.3311 | -12.6496
7.0443 | -34.3316 | -12.6601
7.0504 | -34.3319 | -12.6707
7.0565 | -34.3323 | -12.6811
7.0626 | -34.3327 | -12.6912
7.0687 | -34.333 | -12.701
7.0748 | -34.3334 | -12.7111
7.0809 | -34.3338 | -12.7214
7.087 | -34.3342 | -12.7317
7.0932 | -34.3345 | -12.7418
7.0993 | -34.3349 | -12.7521
7.1054 | -34.3353 | -12.7621
7.1115 | -34.3356 | -12.7722
7.1176 | -34.3359 | -12.7822
7.1237 | -34.3362 | -12.7924
7.1298 | -34.3366 | -12.8021
7.1359 | -34.3369 | -12.812
7.142 | -34.3372 | -12.8222
7.1481 | -34.3376 | -12.8321
7.1542 | -34.3379 | -12.8421
7.1603 | -34.3383 | -12.8521
7.1664 | -34.3387 | -12.8623
7.1725 | -34.3391 | -12.8722
7.1786 | -34.3395 | -12.8825
7.1847 | -34.3399 | -12.8924
7.1908 | -34.3402 | -12.9027
7.1969 | -34.3406 | -12.9128
7.203 | -34.3409 | -12.9229
7.2091 | -34.3413 | -12.9328
7.2152 | -34.3416 | -12.9429
7.2213 | -34.342 | -12.9528
7.2274 | -34.3424 | -12.9628
7.2335 | -34.3427 | -12.9729
7.2397 | -34.3431 | -12.983
7.2458 | -34.3434 | -12.9926
7.2519 | -34.3438 | -13.0027
7.258 | -34.3442 | -13.0129
7.2641 | -34.3445 | -13.0229
7.2702 | -34.3449 | -13.0333
7.2763 | -34.3453 | -13.0431
7.2824 | -34.3457 | -13.0531
7.2885 | -34.3461 | -13.063
7.2946 | -34.3465 | -13.0733
7.3007 | -34.3467 | -13.0833
7.3068 | -34.3471 | -13.0936
7.3129 | -34.3474 | -13.1034
7.319 | -34.3478 | -13.1135
7.3251 | -34.3482 | -13.1232
7.3312 | -34.3486 | -13.1329
7.3373 | -34.349 | -13.1427
7.3434 | -34.3494 | -13.1524
7.3495 | -34.3498 | -13.1619
7.3556 | -34.3501 | -13.1723
7.3617 | -34.3504 | -13.1828
7.3678 | -34.3509 | -13.1929
7.3739 | -34.3512 | -13.2025
7.3801 | -34.3515 | -13.2123
7.3862 | -34.3519 | -13.2223
7.3923 | -34.3523 | -13.2322
7.3984 | -34.3526 | -13.242
7.4045 | -34.353 | -13.2519
7.4106 | -34.3534 | -13.2617
7.4167 | -34.3537 | -13.2718
7.4228 | -34.3541 | -13.2817
7.4289 | -34.3544 | -13.2914
7.435 | -34.3549 | -13.301
7.4411 | -34.3552 | -13.3107
7.4472 | -34.3556 | -13.3207
7.4533 | -34.356 | -13.3307
7.4594 | -34.3563 | -13.3409
7.4655 | -34.3567 | -13.3508
7.4716 | -34.357 | -13.3609
7.4777 | -34.3574 | -13.3703
7.4838 | -34.3578 | -13.3805
7.4899 | -34.3582 | -13.3904
7.496 | -34.3586 | -13.4003
7.5021 | -34.3589 | -13.4101
7.5082 | -34.3593 | -13.4201
7.5143 | -34.3597 | -13.4299
7.5204 | -34.36 | -13.4395
7.5266 | -34.3603 | -13.4494
7.5327 | -34.3607 | -13.4592
7.5388 | -34.361 | -13.469
7.5449 | -34.3613 | -13.4791
7.551 | -34.3617 | -13.4892
7.5571 | -34.3621 | -13.499
7.5632 | -34.3624 | -13.5088
7.5693 | -34.3628 | -13.5188
7.5754 | -34.3632 | -13.5288
7.5815 | -34.3636 | -13.5383
7.5876 | -34.3639 | -13.5481
7.5937 | -34.3642 | -13.5579
7.5998 | -34.3646 | -13.5678
7.6059 | -34.3649 | -13.5776
7.612 | -34.3653 | -13.5874
7.6181 | -34.3656 | -13.5973
7.6242 | -34.366 | -13.6072
7.6303 | -34.3664 | -13.6168
7.6364 | -34.3668 | -13.6269
7.6425 | -34.3671 | -13.6371
7.6486 | -34.3675 | -13.647
7.6547 | -34.3679 | -13.657
7.6608 | -34.3682 | -13.6668
7.667 | -34.3686 | -13.6761
7.6731 | -34.3689 | -13.6862
7.6792 | -34.3692 | -13.6962
7.6853 | -34.3695 | -13.7059
7.6914 | -34.3698 | -13.7159
7.6975 | -34.3702 | -13.7259
7.7036 | -34.3705 | -13.7359
7.7097 | -34.3709 | -13.7457
7.7158 | -34.3712 | -13.7554
7.7219 | -34.3715 | -13.7652
7.728 | -34.3719 | -13.7752
7.7341 | -34.3722 | -13.7851
7.7402 | -34.3726 | -13.7948
7.7463 | -34.3729 | -13.8047
7.7524 | -34.3732 | -13.8149
7.7585 | -34.3735 | -13.8249
7.7646 | -34.3739 | -13.8344
7.7707 | -34.3742 | -13.8441
7.7768 | -34.3744 | -13.8538
7.7829 | -34.3747 | -13.8638
7.789 | -34.3751 | -13.8733
7.7951 | -34.3754 | -13.8836
7.8012 | -34.3758 | -13.8938
7.8073 | -34.3761 | -13.9035
7.8135 | -34.3764 | -13.9135
7.8196 | -34.3767 | -13.9233
7.8257 | -34.3771 | -13.9335
7.8318 | -34.3774 | -13.9438
7.8379 | -34.3778 | -13.9532
7.844 | -34.378 | -13.9626
7.8501 | -34.3783 | -13.9727
7.8562 | -34.3786 | -13.9832
7.8623 | -34.379 | -13.9931
7.8684 | -34.3793 | -14.0027
7.8745 | -34.3796 | -14.0124
7.8806 | -34.3799 | -14.0218
7.8867 | -34.3803 | -14.0315
7.8928 | -34.3806 | -14.0412
7.8989 | -34.381 | -14.0506
7.905 | -34.3814 | -14.0608
7.9111 | -34.3817 | -14.0707
7.9172 | -34.382 | -14.0803
7.9233 | -34.3823 | -14.0902
7.9294 | -34.3827 | -14.1002
7.9355 | -34.383 | -14.1095
7.9416 | -34.3833 | -14.1193
7.9477 | -34.3836 | -14.1287
7.9539 | -34.3839 | -14.1381
7.96 | -34.3843 | -14.1482
7.9661 | -34.3846 | -14.1581
7.9722 | -34.385 | -14.168
7.9783 | -34.3853 | -14.1778
7.9844 | -34.3857 | -14.188
7.9905 | -34.386 | -14.1977
7.9966 | -34.3862 | -14.2075
8.0027 | -34.3866 | -14.2177
8.0088 | -34.3869 | -14.2271
8.0149 | -34.3872 | -14.2367
8.021 | -34.3875 | -14.2465
8.0271 | -34.3879 | -14.2568
8.0332 | -34.3881 | -14.2665
8.0393 | -34.3885 | -14.2765
8.0454 | -34.3888 | -14.2865
8.0515 | -34.3892 | -14.2963
8.0576 | -34.3894 | -14.3064
8.0637 | -34.3897 | -14.3157
8.0698 | -34.3901 | -14.3251
8.0759 | -34.3904 | -14.3346
8.082 | -34.3907 | -14.3446
8.0881 | -34.391 | -14.3543
8.0942 | -34.3914 | -14.3645
8.1004 | -34.3917 | -14.3747
8.1065 | -34.392 | -14.3844
8.1126 | -34.3924 | -14.3941
8.1187 | -34.3927 | -14.404
8.1248 | -34.393 | -14.4141
8.1309 | -34.3933 | -14.4238
8.137 | -34.3936 | -14.4335
8.1431 | -34.394 | -14.4435
8.1492 | -34.3943 | -14.4534
8.1553 | -34.3947 | -14.4632
8.1614 | -34.395 | -14.4733
8.1675 | -34.3953 | -14.483
8.1736 | -34.3956 | -14.4931
8.1797 | -34.3959 | -14.5031
8.1858 | -34.3962 | -14.5128
8.1919 | -34.3965 | -14.5226
8.198 | -34.3968 | -14.532
8.2041 | -34.3972 | -14.5422
8.2102 | -34.3975 | -14.552
8.2163 | -34.3979 | -14.5618
8.2224 | -34.3982 | -14.5718
8.2285 | -34.3985 | -14.5814
8.2346 | -34.3988 | -14.5912
8.2408 | -34.3992 | -14.6011
8.2469 | -34.3995 | -14.6108
8.253 | -34.3998 | -14.6205
8.2591 | -34.4001 | -14.6301
8.2652 | -34.4005 | -14.6397
8.2713 | -34.4009 | -14.6496
8.2774 | -34.4012 | -14.6595
8.2835 | -34.4016 | -14.6697
8.2896 | -34.4019 | -14.679
8.2957 | -34.4022 | -14.6888
8.3018 | -34.4025 | -14.6985
8.3079 | -34.4029 | -14.7083
8.314 | -34.4032 | -14.718
8.3201 | -34.4036 | -14.7277
8.3262 | -34.4039 | -14.7374
8.3323 | -34.4042 | -14.7472
8.3384 | -34.4045 | -14.7574
8.3445 | -34.4049 | -14.7667
8.3506 | -34.4052 | -14.7768
8.3567 | -34.4055 | -14.7868
8.3628 | -34.4058 | -14.7966
8.3689 | -34.4062 | -14.8064
8.375 | -34.4065 | -14.8162
8.3812 | -34.4068 | -14.826
8.3873 | -34.4071 | -14.8358
8.3934 | -34.4075 | -14.8456
8.3995 | -34.4078 | -14.8555
8.4056 | -34.4081 | -14.8654
8.4117 | -34.4085 | -14.8755
8.4178 | -34.4088 | -14.8849
8.4239 | -34.4091 | -14.8944
8.43 | -34.4095 | -14.9042
8.4361 | -34.4098 | -14.9144
8.4422 | -34.4101 | -14.9241
8.4483 | -34.4104 | -14.9339
8.4544 | -34.4108 | -14.9439
8.4605 | -34.4111 | -14.9536
8.4666 | -34.4114 | -14.9636
8.4727 | -34.4117 | -14.9733
8.4788 | -34.412 | -14.9829
8.4849 | -34.4123 | -14.993
8.491 | -34.4127 | -15.0026
8.4971 | -34.413 | -15.0126
8.5032 | -34.4133 | -15.0224
8.5093 | -34.4137 | -15.0321
8.5154 | -34.414 | -15.0417
8.5215 | -34.4143 | -15.0516
8.5277 | -34.4146 | -15.0615
8.5338 | -34.4149 | -15.071
8.5399 | -34.4153 | -15.0811
8.546 | -34.4156 | -15.0906
8.5521 | -34.4159 | -15.1005
8.5582 | -34.4163 | -15.1104
8.5643 | -34.4166 | -15.12
8.5704 | -34.4169 | -15.1295
8.5765 | -34.4172 | -15.1392
8.5826 | -34.4177 | -15.149
8.5887 | -34.418 | -15.1586
8.5948 | -34.4183 | -15.168
8.6009 | -34.4186 | -15.1777
8.607 | -34.4189 | -15.1877
8.6131 | -34.4192 | -15.1978
8.6192 | -34.4196 | -15.2072
8.6253 | -34.4199 | -15.2171
8.6314 | -34.4203 | -15.2268
8.6375 | -34.4206 | -15.2368
8.6436 | -34.421 | -15.2466
8.6497 | -34.4213 | -15.2565
8.6558 | -34.4216 | -15.266
8.6619 | -34.4219 | -15.276
8.6681 | -34.4222 | -15.2857
8.6742 | -34.4226 | -15.2956
8.6803 | -34.4229 | -15.3057
8.6864 | -34.4233 | -15.3157
8.6925 | -34.4237 | -15.3255
8.6986 | -34.424 | -15.3356
8.7047 | -34.4244 | -15.3455
8.7108 | -34.4247 | -15.3553
8.7169 | -34.425 | -15.3649
8.723 | -34.4254 | -15.3747
8.7291 | -34.4257 | -15.3844
8.7352 | -34.4261 | -15.3944
8.7413 | -34.4264 | -15.4041
8.7474 | -34.4268 | -15.4136
8.7535 | -34.4272 | -15.4235
8.7596 | -34.4275 | -15.4338
8.7657 | -34.4278 | -15.4436
8.7718 | -34.4282 | -15.4531
8.7779 | -34.4285 | -15.4634
8.784 | -34.4288 | -15.473
8.7901 | -34.4291 | -15.4833
8.7962 | -34.4295 | -15.4932
8.8023 | -34.4299 | -15.5026
8.8084 | -34.4302 | -15.5126
8.8146 | -34.4305 | -15.5223
8.8207 | -34.4309 | -15.5321
8.8268 | -34.4313 | -15.5418
8.8329 | -34.4316 | -15.5517
8.839 | -34.432 | -15.5616
8.8451 | -34.4323 | -15.5712
8.8512 | -34.4327 | -15.5811
8.8573 | -34.433 | -15.5907
8.8634 | -34.4335 | -15.6004
8.8695 | -34.4338 | -15.6103
8.8756 | -34.4341 | -15.6199
8.8817 | -34.4345 | -15.6299
8.8878 | -34.4348 | -15.6399
8.8939 | -34.4351 | -15.6494
8.9 | -34.4354 | -15.6589
8.9061 | -34.4358 | -15.6687
8.9122 | -34.4362 | -15.6784
8.9183 | -34.4365 | -15.688
8.9244 | -34.4369 | -15.6976
8.9305 | -34.4373 | -15.7075
8.9366 | -34.4376 | -15.7171
8.9427 | -34.438 | -15.727
8.9488 | -34.4384 | -15.7367
8.955 | -34.4387 | -15.7465
8.9611 | -34.4391 | -15.7564
8.9672 | -34.4395 | -15.7665
8.9733 | -34.4398 | -15.7762
8.9794 | -34.4402 | -15.7861
8.9855 | -34.4406 | -15.7958
8.9916 | -34.4409 | -15.8054
8.9977 | -34.4413 | -15.8151
9.0038 | -34.4417 | -15.8251
9.0099 | -34.4421 | -15.835
9.016 | -34.4425 | -15.8447
9.0221 | -34.4428 | -15.8547
9.0282 | -34.4432 | -15.8647
9.0343 | -34.4436 | -15.8747
9.0404 | -34.444 | -15.8844
9.0465 | -34.4444 | -15.8945
9.0526 | -34.4448 | -15.9042
9.0587 | -34.4452 | -15.914
9.0648 | -34.4456 | -15.9237
9.0709 | -34.4459 | -15.9332
9.077 | -34.4463 | -15.9436
9.0831 | -34.4466 | -15.9534
9.0892 | -34.447 | -15.9631
9.0953 | -34.4473 | -15.973
9.1015 | -34.4477 | -15.9828
9.1076 | -34.4481 | -15.9927
9.1137 | -34.4485 | -16.0024
9.1198 | -34.4489 | -16.0118
9.1259 | -34.4493 | -16.0217
9.132 | -34.4497 | -16.0318
9.1381 | -34.4501 | -16.0414
9.1442 | -34.4505 | -16.0508
9.1503 | -34.4509 | -16.0607
9.1564 | -34.4512 | -16.0705
9.1625 | -34.4516 | -16.0804
9.1686 | -34.452 | -16.0902
9.1747 | -34.4524 | -16.1001
9.1808 | -34.4528 | -16.1101
9.1869 | -34.4531 | -16.1202
9.193 | -34.4535 | -16.1302
9.1991 | -34.4539 | -16.1397
9.2052 | -34.4543 | -16.1494
9.2113 | -34.4547 | -16.1594
9.2174 | -34.4551 | -16.1699
9.2235 | -34.4555 | -16.1794
9.2296 | -34.4559 | -16.1893
9.2357 | -34.4563 | -16.1991
9.2419 | -34.4567 | -16.2089
9.248 | -34.4571 | -16.219
9.2541 | -34.4575 | -16.2289
9.2602 | -34.4578 | -16.2389
9.2663 | -34.4582 | -16.2485
9.2724 | -34.4586 | -16.2584
9.2785 | -34.459 | -16.2682
9.2846 | -34.4594 | -16.2781
9.2907 | -34.4598 | -16.2879
9.2968 | -34.4601 | -16.2978
9.3029 | -34.4605 | -16.3077
9.309 | -34.4609 | -16.3175
9.3151 | -34.4613 | -16.3272
9.3212 | -34.4618 | -16.3372
9.3273 | -34.4622 | -16.3469
9.3334 | -34.4626 | -16.3565
9.3395 | -34.4631 | -16.3662
9.3456 | -34.4635 | -16.3758
9.3517 | -34.4639 | -16.3853
9.3578 | -34.4643 | -16.3949
9.3639 | -34.4647 | -16.4049
9.37 | -34.4651 | -16.4147
9.3761 | -34.4655 | -16.4244
9.3822 | -34.4659 | -16.4341
9.3884 | -34.4664 | -16.4439
9.3945 | -34.4668 | -16.4538
9.4006 | -34.4671 | -16.4637
9.4067 | -34.4675 | -16.4736
9.4128 | -34.468 | -16.4834
9.4189 | -34.4684 | -16.4933
9.425 | -34.4688 | -16.503
9.4311 | -34.4692 | -16.5128
9.4372 | -34.4696 | -16.5226
9.4433 | -34.47 | -16.5325
9.4494 | -34.4704 | -16.5423
9.4555 | -34.4708 | -16.552
9.4616 | -34.4713 | -16.5618
9.4677 | -34.4717 | -16.5715
9.4738 | -34.4721 | -16.5816
9.4799 | -34.4725 | -16.5913
9.486 | -34.4728 | -16.6015
9.4921 | -34.4733 | -16.6117
9.4982 | -34.4737 | -16.621
9.5043 | -34.4741 | -16.6306
9.5104 | -34.4745 | -16.6402
9.5165 | -34.4749 | -16.65
9.5226 | -34.4753 | -16.66
9.5288 | -34.4758 | -16.6698
9.5349 | -34.4762 | -16.6796
9.541 | -34.4766 | -16.6893
9.5471 | -34.477 | -16.6992
9.5532 | -34.4774 | -16.7089
9.5593 | -34.4779 | -16.7187
9.5654 | -34.4783 | -16.7286
9.5715 | -34.4787 | -16.7386
9.5776 | -34.4791 | -16.7487
9.5837 | -34.4795 | -16.7582
9.5898 | -34.4799 | -16.7682
9.5959 | -34.4803 | -16.7779
9.602 | -34.4807 | -16.7878
9.6081 | -34.4811 | -16.7977
9.6142 | -34.4816 | -16.8074
9.6203 | -34.4819 | -16.8173
9.6264 | -34.4823 | -16.8267
9.6325 | -34.4826 | -16.8366
9.6386 | -34.4831 | -16.8464
9.6447 | -34.4835 | -16.8561
9.6508 | -34.4839 | -16.866
9.6569 | -34.4843 | -16.8757
9.663 | -34.4848 | -16.8855
9.6691 | -34.4852 | -16.8954
9.6753 | -34.4857 | -16.9052
9.6814 | -34.4861 | -16.9147
9.6875 | -34.4866 | -16.9242
9.6936 | -34.487 | -16.9339
9.6997 | -34.4874 | -16.9438
9.7058 | -34.4878 | -16.9536
9.7119 | -34.4882 | -16.9633
9.718 | -34.4886 | -16.9734
9.7241 | -34.489 | -16.9832
9.7302 | -34.4894 | -16.993
9.7363 | -34.4899 | -17.0026
9.7424 | -34.4903 | -17.0123
9.7485 | -34.4907 | -17.0219
9.7546 | -34.4912 | -17.0314
9.7607 | -34.4915 | -17.0414
9.7668 | -34.492 | -17.0516
9.7729 | -34.4924 | -17.0611
9.779 | -34.4929 | -17.0709
9.7851 | -34.4933 | -17.0808
9.7912 | -34.4937 | -17.0911
9.7973 | -34.4941 | -17.1011
9.8034 | -34.4945 | -17.1111
9.8095 | -34.4949 | -17.1207
9.8157 | -34.4953 | -17.1305
9.8218 | -34.4958 | -17.1399
9.8279 | -34.4962 | -17.1498
9.834 | -34.4966 | -17.1594
9.8401 | -34.4971 | -17.169
9.8462 | -34.4975 | -17.1785
9.8523 | -34.498 | -17.1882
9.8584 | -34.4984 | -17.1982
9.8645 | -34.4988 | -17.208
9.8706 | -34.4992 | -17.2176
9.8767 | -34.4996 | -17.227
9.8828 | -34.5001 | -17.2364
9.8889 | -34.5005 | -17.2464
9.895 | -34.5009 | -17.2558
9.9011 | -34.5013 | -17.265
9.9072 | -34.5017 | -17.2744
9.9133 | -34.5022 | -17.2839
9.9194 | -34.5026 | -17.2934
9.9255 | -34.503 | -17.3033
9.9316 | -34.5035 | -17.3128
9.9377 | -34.5039 | -17.3223
9.9438 | -34.5044 | -17.3319
9.9499 | -34.5048 | -17.3417
9.956 | -34.5052 | -17.3512
9.9622 | -34.5056 | -17.3608
9.9683 | -34.5061 | -17.3705
9.9744 | -34.5065 | -17.3801
9.9805 | -34.5069 | -17.3899
9.9866 | -34.5074 | -17.3992
9.9927 | -34.5078 | -17.409
9.9988 | -34.5082 | -17.4185
10.0049 | -34.5086 | -17.4282
10.011 | -34.5091 | -17.4377
10.0171 | -34.5095 | -17.4474
10.0232 | -34.51 | -17.4567
10.0293 | -34.5105 | -17.4662
10.0354 | -34.511 | -17.4755
10.0415 | -34.5114 | -17.4851
10.0476 | -34.5118 | -17.495
10.0537 | -34.5122 | -17.5049
10.0598 | -34.5126 | -17.5147
10.0659 | -34.513 | -17.5243
10.072 | -34.5135 | -17.5338
10.0781 | -34.5139 | -17.5434
10.0842 | -34.5144 | -17.5532
10.0903 | -34.5148 | -17.5631
10.0964 | -34.5153 | -17.5724
10.1026 | -34.5158 | -17.582
10.1087 | -34.5162 | -17.5916
10.1148 | -34.5167 | -17.6013
10.1209 | -34.5171 | -17.6108
10.127 | -34.5175 | -17.6201
10.1331 | -34.518 | -17.6296
10.1392 | -34.5184 | -17.6391
10.1453 | -34.5189 | -17.6488
10.1514 | -34.5193 | -17.6581
10.1575 | -34.5197 | -17.6676
10.1636 | -34.5202 | -17.6764
10.1697 | -34.5206 | -17.686
10.1758 | -34.5211 | -17.6954
10.1819 | -34.5215 | -17.7051
10.188 | -34.5219 | -17.7144
10.1941 | -34.5223 | -17.7241
10.2002 | -34.5227 | -17.733
10.2063 | -34.5232 | -17.7428
10.2124 | -34.5236 | -17.7524
10.2185 | -34.524 | -17.7619
10.2246 | -34.5245 | -17.7714
10.2307 | -34.5249 | -17.7811
10.2368 | -34.5254 | -17.7897
10.2429 | -34.5258 | -17.7989
10.2491 | -34.5262 | -17.8082
10.2552 | -34.5267 | -17.8181
10.2613 | -34.5272 | -17.8279
10.2674 | -34.5275 | -17.8376
10.2735 | -34.5279 | -17.8473
10.2796 | -34.5284 | -17.8569
10.2857 | -34.5288 | -17.8665
10.2918 | -34.5292 | -17.876
10.2979 | -34.5297 | -17.8857
10.304 | -34.5302 | -17.8954
10.3101 | -34.5306 | -17.9049
10.3162 | -34.5311 | -17.9144
10.3223 | -34.5315 | -17.9239
10.3284 | -34.5319 | -17.9333
10.3345 | -34.5324 | -17.9426
10.3406 | -34.5328 | -17.952
10.3467 | -34.5332 | -17.9614
10.3528 | -34.5336 | -17.9706
10.3589 | -34.534 | -17.9803
10.365 | -34.5345 | -17.99
10.3711 | -34.5349 | -17.9998
10.3772 | -34.5353 | -18.0093
10.3833 | -34.5357 | -18.0187
10.3895 | -34.5361 | -18.028
10.3956 | -34.5366 | -18.0373
10.4017 | -34.537 | -18.0467
10.4078 | -34.5374 | -18.0562
10.4139 | -34.5379 | -18.0656
10.42 | -34.5384 | -18.0751
10.4261 | -34.5388 | -18.0845
10.4322 | -34.5393 | -18.0938
10.4383 | -34.5397 | -18.1032
10.4444 | -34.5401 | -18.1122
10.4505 | -34.5406 | -18.1217
10.4566 | -34.541 | -18.1311
10.4627 | -34.5414 | -18.1405
10.4688 | -34.5419 | -18.1495
10.4749 | -34.5423 | -18.1587
10.481 | -34.5428 | -18.1682
10.4871 | -34.5432 | -18.1777
10.4932 | -34.5437 | -18.1871
10.4993 | -34.5442 | -18.1964
10.5054 | -34.5446 | -18.2059
10.5115 | -34.545 | -18.2153
10.5176 | -34.5455 | -18.2244
10.5237 | -34.546 | -18.2337
10.5298 | -34.5463 | -18.2428
10.536 | -34.5468 | -18.2524
10.5421 | -34.5473 | -18.2614
10.5482 | -34.5477 | -18.2711
10.5543 | -34.5482 | -18.2804
10.5604 | -34.5486 | -18.2894
10.5665 | -34.5491 | -18.2986
10.5726 | -34.5495 | -18.3078
10.5787 | -34.5499 | -18.3173
10.5848 | -34.5504 | -18.3262
10.5909 | -34.5508 | -18.3357
10.597 | -34.5512 | -18.345
10.6031 | -34.5517 | -18.3545
10.6092 | -34.5522 | -18.3641
10.6153 | -34.5526 | -18.3737
10.6214 | -34.5531 | -18.3829
10.6275 | -34.5535 | -18.3921
10.6336 | -34.554 | -18.4018
10.6397 | -34.5544 | -18.4111
10.6458 | -34.5548 | -18.4201
10.6519 | -34.5552 | -18.4294
10.658 | -34.5556 | -18.4385
10.6641 | -34.5561 | -18.4478
10.6702 | -34.5565 | -18.4574
10.6764 | -34.557 | -18.4667
10.6825 | -34.5575 | -18.4761
10.6886 | -34.558 | -18.4855
10.6947 | -34.5584 | -18.4946
10.7008 | -34.5588 | -18.5039
10.7069 | -34.5593 | -18.5133
10.713 | -34.5598 | -18.5225
10.7191 | -34.5602 | -18.5319
10.7252 | -34.5607 | -18.541
10.7313 | -34.5612 | -18.5503
10.7374 | -34.5616 | -18.5593
10.7435 | -34.562 | -18.5689
10.7496 | -34.5625 | -18.5779
10.7557 | -34.5629 | -18.5872
10.7618 | -34.5634 | -18.5965
10.7679 | -34.5638 | -18.6057
10.774 | -34.5643 | -18.615
10.7801 | -34.5647 | -18.6247
10.7862 | -34.5652 | -18.6338
10.7923 | -34.5656 | -18.6432
10.7984 | -34.5661 | -18.6523
10.8045 | -34.5665 | -18.6618
10.8106 | -34.567 | -18.671
10.8168 | -34.5674 | -18.6803
10.8229 | -34.5678 | -18.6897
10.829 | -34.5683 | -18.699
10.8351 | -34.5687 | -18.7083
10.8412 | -34.5692 | -18.7175
10.8473 | -34.5696 | -18.7264
10.8534 | -34.57 | -18.7359
10.8595 | -34.5705 | -18.7449
10.8656 | -34.571 | -18.754
10.8717 | -34.5714 | -18.7629
10.8778 | -34.5719 | -18.772
10.8839 | -34.5723 | -18.7812
10.89 | -34.5728 | -18.7902
10.8961 | -34.5732 | -18.7995
10.9022 | -34.5737 | -18.8086
10.9083 | -34.5742 | -18.8181
10.9144 | -34.5746 | -18.8271
10.9205 | -34.5751 | -18.8364
10.9266 | -34.5756 | -18.8459
10.9327 | -34.5761 | -18.8551
10.9388 | -34.5765 | -18.8643
10.9449 | -34.577 | -18.8735
10.951 | -34.5774 | -18.8827
10.9571 | -34.5778 | -18.892
10.9633 | -34.5783 | -18.9016
10.9694 | -34.5788 | -18.9111
10.9755 | -34.5793 | -18.9203
10.9816 | -34.5797 | -18.9296
10.9877 | -34.5802 | -18.9388
10.9938 | -34.5806 | -18.9482
10.9999 | -34.5811 | -18.9573
11.006 | -34.5816 | -18.9664
11.0121 | -34.582 | -18.9759
11.0182 | -34.5825 | -18.9848
11.0243 | -34.5829 | -18.9941
11.0304 | -34.5833 | -19.0032
11.0365 | -34.5838 | -19.0123
11.0426 | -34.5843 | -19.0213
11.0487 | -34.5847 | -19.0306
11.0548 | -34.5852 | -19.0397
11.0609 | -34.5857 | -19.0492
11.067 | -34.5861 | -19.0586
11.0731 | -34.5866 | -19.0678
11.0792 | -34.5871 | -19.077
11.0853 | -34.5875 | -19.0859
11.0914 | -34.588 | -19.095
11.0975 | -34.5885 | -19.1044
11.1037 | -34.5889 | -19.1133
11.1098 | -34.5893 | -19.1228
11.1159 | -34.5898 | -19.132
11.122 | -34.5903 | -19.1415
11.1281 | -34.5907 | -19.1504
11.1342 | -34.5911 | -19.1593
11.1403 | -34.5916 | -19.1688
11.1464 | -34.5921 | -19.178
11.1525 | -34.5925 | -19.1872
11.1586 | -34.593 | -19.1969
11.1647 | -34.5935 | -19.2059
11.1708 | -34.5939 | -19.215
11.1769 | -34.5943 | -19.2236
11.183 | -34.5948 | -19.2327
11.1891 | -34.5953 | -19.2423
11.1952 | -34.5958 | -19.2518
11.2013 | -34.5962 | -19.2609
11.2074 | -34.5966 | -19.2698
11.2135 | -34.597 | -19.279
11.2196 | -34.5975 | -19.2881
11.2257 | -34.5979 | -19.2971
11.2318 | -34.5983 | -19.3065
11.2379 | -34.5987 | -19.3156
11.244 | -34.5991 | -19.3247
11.2502 | -34.5996 | -19.3338
11.2563 | -34.6001 | -19.343
11.2624 | -34.6005 | -19.3522
11.2685 | -34.6009 | -19.3612
11.2746 | -34.6014 | -19.37
11.2807 | -34.6018 | -19.3796
11.2868 | -34.6022 | -19.3891
11.2929 | -34.6028 | -19.3979
11.299 | -34.6032 | -19.4073
11.3051 | -34.6036 | -19.4164
11.3112 | -34.604 | -19.4257
11.3173 | -34.6045 | -19.4349
11.3234 | -34.605 | -19.4441
11.3295 | -34.6054 | -19.4531
11.3356 | -34.6059 | -19.4622
11.3417 | -34.6063 | -19.4715
11.3478 | -34.6068 | -19.4806
11.3539 | -34.6072 | -19.4897
11.36 | -34.6076 | -19.4987
11.3661 | -34.6081 | -19.5077
11.3722 | -34.6085 | -19.5167
11.3783 | -34.609 | -19.5258
11.3844 | -34.6094 | -19.535
11.3906 | -34.6098 | -19.5444
11.3967 | -34.6103 | -19.5534
11.4028 | -34.6107 | -19.5627
11.4089 | -34.6112 | -19.5717
11.415 | -34.6116 | -19.5807
11.4211 | -34.612 | -19.5896
11.4272 | -34.6125 | -19.5988
11.4333 | -34.6129 | -19.608
11.4394 | -34.6134 | -19.6175
11.4455 | -34.6138 | -19.6266
11.4516 | -34.6142 | -19.6356
11.4577 | -34.6147 | -19.645
11.4638 | -34.6152 | -19.6537
11.4699 | -34.6156 | -19.6632
11.476 | -34.616 | -19.6728
11.4821 | -34.6165 | -19.6816
11.4882 | -34.617 | -19.691
11.4943 | -34.6174 | -19.7
11.5004 | -34.6179 | -19.7092
11.5065 | -34.6184 | -19.7183
11.5126 | -34.6188 | -19.7271
11.5187 | -34.6193 | -19.7362
11.5248 | -34.6198 | -19.7454
11.5309 | -34.6202 | -19.7545
11.5371 | -34.6206 | -19.7635
11.5432 | -34.621 | -19.7728
11.5493 | -34.6215 | -19.7821
11.5554 | -34.6219 | -19.791
11.5615 | -34.6224 | -19.8004
11.5676 | -34.6229 | -19.8097
11.5737 | -34.6233 | -19.8185
11.5798 | -34.6238 | -19.8274
11.5859 | -34.6243 | -19.8364
11.592 | -34.6247 | -19.8455
11.5981 | -34.6252 | -19.8545
11.6042 | -34.6256 | -19.8638
11.6103 | -34.626 | -19.8727
11.6164 | -34.6265 | -19.8816
11.6225 | -34.6269 | -19.8907
11.6286 | -34.6274 | -19.8999
11.6347 | -34.6278 | -19.9093
11.6408 | -34.6283 | -19.9183
11.6469 | -34.6288 | -19.9276
11.653 | -34.6293 | -19.9365
11.6591 | -34.6298 | -19.9456
11.6652 | -34.6302 | -19.9548
11.6713 | -34.6306 | -19.9639
11.6775 | -34.6311 | -19.9728
11.6836 | -34.6316 | -19.982
11.6897 | -34.632 | -19.9909
11.6958 | -34.6325 | -20.0003
11.7019 | -34.633 | -20.0095
11.708 | -34.6334 | -20.0188
11.7141 | -34.6339 | -20.0279
11.7202 | -34.6344 | -20.0367
11.7263 | -34.6348 | -20.0459
11.7324 | -34.6353 | -20.0552
11.7385 | -34.6357 | -20.0644
11.7446 | -34.6361 | -20.0737
11.7507 | -34.6366 | -20.0827
11.7568 | -34.6371 | -20.0919
11.7629 | -34.6375 | -20.101
11.769 | -34.638 | -20.1101
11.7751 | -34.6385 | -20.1189
11.7812 | -34.6389 | -20.1278
11.7873 | -34.6394 | -20.1372
11.7934 | -34.6399 | -20.1463
11.7995 | -34.6403 | -20.155
11.8056 | -34.6407 | -20.1645
11.8117 | -34.6411 | -20.1736
11.8178 | -34.6416 | -20.1827
11.824 | -34.6422 | -20.1915
11.8301 | -34.6427 | -20.2002
11.8362 | -34.6431 | -20.2089
11.8423 | -34.6435 | -20.2181
11.8484 | -34.644 | -20.2271
11.8545 | -34.6444 | -20.2365
11.8606 | -34.6449 | -20.2456
11.8667 | -34.6454 | -20.2547
11.8728 | -34.6458 | -20.2638
11.8789 | -34.6463 | -20.2732
11.885 | -34.6467 | -20.2822
11.8911 | -34.6472 | -20.2915
11.8972 | -34.6476 | -20.3002
11.9033 | -34.6481 | -20.3089
11.9094 | -34.6485 | -20.3176
11.9155 | -34.6489 | -20.3266
11.9216 | -34.6494 | -20.3355
11.9277 | -34.6499 | -20.3444
11.9338 | -34.6503 | -20.3535
11.9399 | -34.6508 | -20.3629
11.946 | -34.6513 | -20.3719
11.9521 | -34.6517 | -20.3809
11.9582 | -34.6522 | -20.3897
11.9644 | -34.6527 | -20.3986
11.9705 | -34.6531 | -20.4078
11.9766 | -34.6536 | -20.4168
11.9827 | -34.654 | -20.4257
11.9888 | -34.6545 | -20.4345
11.9949 | -34.6549 | -20.4435
12.001 | -34.6554 | -20.4526
12.0071 | -34.6559 | -20.4614
12.0132 | -34.6563 | -20.4701
12.0193 | -34.6568 | -20.4794
12.0254 | -34.6572 | -20.4882
12.0315 | -34.6577 | -20.4971
12.0376 | -34.6582 | -20.5059
12.0437 | -34.6586 | -20.5147
12.0498 | -34.6591 | -20.5239
12.0559 | -34.6595 | -20.5329
12.062 | -34.66 | -20.5419
12.0681 | -34.6605 | -20.5508
12.0742 | -34.661 | -20.56
12.0803 | -34.6614 | -20.5688
12.0864 | -34.6619 | -20.5778
12.0925 | -34.6624 | -20.5868
12.0986 | -34.6629 | -20.5955
12.1047 | -34.6633 | -20.6045
12.1109 | -34.6638 | -20.6136
12.117 | -34.6643 | -20.6225
12.1231 | -34.6648 | -20.6314
12.1292 | -34.6652 | -20.6405
12.1353 | -34.6657 | -20.6492
12.1414 | -34.6662 | -20.6584
12.1475 | -34.6666 | -20.6674
12.1536 | -34.6671 | -20.6765
12.1597 | -34.6676 | -20.6855
12.1658 | -34.668 | -20.6945
12.1719 | -34.6685 | -20.7035
12.178 | -34.669 | -20.7126
12.1841 | -34.6695 | -20.7217
12.1902 | -34.67 | -20.7307
12.1963 | -34.6704 | -20.7396
12.2024 | -34.6709 | -20.7485
12.2085 | -34.6714 | -20.7575
12.2146 | -34.6718 | -20.7663
12.2207 | -34.6723 | -20.7753
12.2268 | -34.6727 | -20.7842
12.2329 | -34.6732 | -20.7932
12.239 | -34.6736 | -20.8019
12.2451 | -34.6741 | -20.8106
12.2513 | -34.6746 | -20.8196
12.2574 | -34.6751 | -20.8289
12.2635 | -34.6756 | -20.8374
12.2696 | -34.676 | -20.8461
12.2757 | -34.6765 | -20.8547
12.2818 | -34.6769 | -20.8634
12.2879 | -34.6773 | -20.8722
12.294 | -34.6778 | -20.8809
12.3001 | -34.6783 | -20.89
12.3062 | -34.6788 | -20.8989
12.3123 | -34.6793 | -20.9078
12.3184 | -34.6798 | -20.9166
12.3245 | -34.6803 | -20.9256
12.3306 | -34.6807 | -20.9342
12.3367 | -34.6812 | -20.943
12.3428 | -34.6817 | -20.9516
12.3489 | -34.6821 | -20.96
12.355 | -34.6826 | -20.9688
12.3611 | -34.683 | -20.9777
12.3672 | -34.6835 | -20.9865
12.3733 | -34.684 | -20.9957
12.3794 | -34.6844 | -21.0041
12.3855 | -34.6849 | -21.0128
12.3916 | -34.6853 | -21.0219
12.3978 | -34.6858 | -21.031
12.4039 | -34.6862 | -21.0398
12.41 | -34.6867 | -21.0489
12.4161 | -34.6872 | -21.0574
12.4222 | -34.6877 | -21.0661
12.4283 | -34.6882 | -21.075
12.4344 | -34.6886 | -21.0838
12.4405 | -34.6891 | -21.0927
12.4466 | -34.6896 | -21.1014
12.4527 | -34.69 | -21.1103
12.4588 | -34.6905 | -21.1191
12.4649 | -34.691 | -21.1282
12.471 | -34.6915 | -21.1375
12.4771 | -34.6919 | -21.1457
12.4832 | -34.6924 | -21.1545
12.4893 | -34.6929 | -21.1631
12.4954 | -34.6934 | -21.1718
12.5015 | -34.6939 | -21.1804
12.5076 | -34.6943 | -21.1892
12.5137 | -34.6948 | -21.1985
12.5198 | -34.6952 | -21.207
12.5259 | -34.6957 | -21.2156
12.532 | -34.6962 | -21.2246
12.5382 | -34.6967 | -21.2332
12.5443 | -34.6972 | -21.2421
12.5504 | -34.6976 | -21.251
12.5565 | -34.6982 | -21.2595
12.5626 | -34.6986 | -21.2686
12.5687 | -34.6991 | -21.2775
12.5748 | -34.6996 | -21.2862
12.5809 | -34.7 | -21.2947
12.587 | -34.7004 | -21.3038
12.5931 | -34.7009 | -21.3126
12.5992 | -34.7014 | -21.3216
12.6053 | -34.7019 | -21.3303
12.6114 | -34.7024 | -21.3391
12.6175 | -34.7028 | -21.348
12.6236 | -34.7033 | -21.3565
12.6297 | -34.7038 | -21.3651
12.6358 | -34.7043 | -21.3736
12.6419 | -34.7048 | -21.3826
12.648 | -34.7053 | -21.3916
12.6541 | -34.7057 | -21.4003
12.6602 | -34.7063 | -21.4087
12.6663 | -34.7067 | -21.4173
12.6724 | -34.7072 | -21.4261
12.6785 | -34.7076 | -21.4351
12.6847 | -34.7081 | -21.4439
12.6908 | -34.7086 | -21.4526
12.6969 | -34.7091 | -21.4613
12.703 | -34.7096 | -21.4701
12.7091 | -34.71 | -21.4789
12.7152 | -34.7105 | -21.4876
12.7213 | -34.711 | -21.4965
12.7274 | -34.7114 | -21.5053
12.7335 | -34.7119 | -21.5135
12.7396 | -34.7124 | -21.5218
12.7457 | -34.7128 | -21.5306
12.7518 | -34.7133 | -21.5393
12.7579 | -34.7137 | -21.5483
12.764 | -34.7142 | -21.5571
12.7701 | -34.7146 | -21.5655
12.7762 | -34.7151 | -21.5742
12.7823 | -34.7155 | -21.583
12.7884 | -34.716 | -21.5917
12.7945 | -34.7164 | -21.6
12.8006 | -34.7169 | -21.6081
12.8067 | -34.7174 | -21.617
12.8128 | -34.7178 | -21.6253
12.8189 | -34.7182 | -21.6337
12.8251 | -34.7186 | -21.6422
12.8312 | -34.7191 | -21.651
12.8373 | -34.7196 | -21.6596
12.8434 | -34.7201 | -21.6682
12.8495 | -34.7205 | -21.6767
12.8556 | -34.7209 | -21.6853
12.8617 | -34.7214 | -21.6937
12.8678 | -34.7218 | -21.7024
12.8739 | -34.7223 | -21.7106
12.88 | -34.7228 | -21.7193
12.8861 | -34.7232 | -21.728
12.8922 | -34.7236 | -21.7366
12.8983 | -34.7242 | -21.7449
12.9044 | -34.7246 | -21.7535
12.9105 | -34.7252 | -21.762
12.9166 | -34.7256 | -21.7709
12.9227 | -34.7261 | -21.78
12.9288 | -34.7266 | -21.7888
12.9349 | -34.7271 | -21.7978
12.941 | -34.7276 | -21.8065
12.9471 | -34.7281 | -21.8151
12.9532 | -34.7286 | -21.8237
12.9593 | -34.7291 | -21.8325
12.9654 | -34.7295 | -21.8413
12.9716 | -34.73 | -21.8499
12.9777 | -34.7304 | -21.8582
12.9838 | -34.7309 | -21.8667
12.9899 | -34.7313 | -21.8754
12.996 | -34.7317 | -21.8841
13.0021 | -34.7322 | -21.8928
13.0082 | -34.7327 | -21.9014
13.0143 | -34.7331 | -21.9096
13.0204 | -34.7335 | -21.9179
13.0265 | -34.734 | -21.9261
13.0326 | -34.7345 | -21.935
13.0387 | -34.7349 | -21.9438
13.0448 | -34.7354 | -21.9524
13.0509 | -34.7359 | -21.961
13.057 | -34.7364 | -21.9692
13.0631 | -34.7369 | -21.9777
13.0692 | -34.7374 | -21.9861
13.0753 | -34.7379 | -21.9946
13.0814 | -34.7383 | -22.0029
13.0875 | -34.7388 | -22.0111
13.0936 | -34.7392 | -22.0195
13.0997 | -34.7397 | -22.0276
13.1058 | -34.7402 | -22.0356
13.112 | -34.7406 | -22.0441
13.1181 | -34.7411 | -22.0527
13.1242 | -34.7415 | -22.0613
13.1303 | -34.742 | -22.0701
13.1364 | -34.7424 | -22.0786
13.1425 | -34.7429 | -22.087
13.1486 | -34.7433 | -22.0954
13.1547 | -34.7438 | -22.1036
13.1608 | -34.7443 | -22.112
13.1669 | -34.7447 | -22.1204
13.173 | -34.7452 | -22.1289
13.1791 | -34.7456 | -22.1375
13.1852 | -34.7461 | -22.1462
13.1913 | -34.7466 | -22.1547
13.1974 | -34.747 | -22.1632
13.2035 | -34.7474 | -22.1716
13.2096 | -34.7479 | -22.1801
13.2157 | -34.7484 | -22.1886
13.2218 | -34.7489 | -22.1972
13.2279 | -34.7494 | -22.2057
13.234 | -34.7499 | -22.2143
13.2401 | -34.7503 | -22.2227
13.2462 | -34.7508 | -22.2311
13.2524 | -34.7512 | -22.2394
13.2585 | -34.7517 | -22.2478
13.2646 | -34.7521 | -22.2565
13.2707 | -34.7526 | -22.2651
13.2768 | -34.7531 | -22.2739
13.2829 | -34.7535 | -22.2824
13.289 | -34.754 | -22.2909
13.2951 | -34.7545 | -22.2991
13.3012 | -34.7549 | -22.3076
13.3073 | -34.7553 | -22.3164
13.3134 | -34.7558 | -22.3248
13.3195 | -34.7562 | -22.3331
13.3256 | -34.7567 | -22.3417
13.3317 | -34.7572 | -22.35
13.3378 | -34.7576 | -22.3585
13.3439 | -34.7581 | -22.367
13.35 | -34.7585 | -22.3755
13.3561 | -34.759 | -22.384
13.3622 | -34.7594 | -22.3925
13.3683 | -34.7599 | -22.4007
13.3744 | -34.7603 | -22.409
13.3805 | -34.7607 | -22.4175
13.3866 | -34.7612 | -22.4257
13.3927 | -34.7616 | -22.4341
13.3989 | -34.7621 | -22.4421
13.405 | -34.7626 | -22.4504
13.4111 | -34.7631 | -22.4588
13.4172 | -34.7635 | -22.4668
13.4233 | -34.7639 | -22.4755
13.4294 | -34.7644 | -22.4835
13.4355 | -34.7649 | -22.4925
13.4416 | -34.7653 | -22.501
13.4477 | -34.7657 | -22.5095
13.4538 | -34.7662 | -22.5178
13.4599 | -34.7667 | -22.5264
13.466 | -34.7671 | -22.535
13.4721 | -34.7676 | -22.5435
13.4782 | -34.768 | -22.5519
13.4843 | -34.7686 | -22.5603
13.4904 | -34.769 | -22.5683
13.4965 | -34.7694 | -22.5771
13.5026 | -34.7699 | -22.5855
13.5087 | -34.7704 | -22.5934
13.5148 | -34.7708 | -22.6019
13.5209 | -34.7712 | -22.6101
13.527 | -34.7717 | -22.6189
13.5331 | -34.7721 | -22.6272
13.5393 | -34.7726 | -22.6356
13.5454 | -34.773 | -22.6441
13.5515 | -34.7735 | -22.6525
13.5576 | -34.7739 | -22.661
13.5637 | -34.7744 | -22.6693
13.5698 | -34.7749 | -22.6779
13.5759 | -34.7752 | -22.6859
13.582 | -34.7757 | -22.6944
13.5881 | -34.7761 | -22.7028
13.5942 | -34.7766 | -22.7113
13.6003 | -34.7771 | -22.7198
13.6064 | -34.7775 | -22.7284
13.6125 | -34.778 | -22.7371
13.6186 | -34.7785 | -22.7455
13.6247 | -34.7789 | -22.7538
13.6308 | -34.7793 | -22.7619
13.6369 | -34.7798 | -22.7703
13.643 | -34.7803 | -22.7786
13.6491 | -34.7807 | -22.7871
13.6552 | -34.7812 | -22.7953
13.6613 | -34.7817 | -22.8039
13.6674 | -34.7821 | -22.8124
13.6735 | -34.7826 | -22.8209
13.6796 | -34.783 | -22.8292
13.6858 | -34.7834 | -22.8377
13.6919 | -34.7839 | -22.8462
13.698 | -34.7844 | -22.8546
13.7041 | -34.7848 | -22.8633
13.7102 | -34.7853 | -22.8718
13.7163 | -34.7858 | -22.8803
13.7224 | -34.7862 | -22.8886
13.7285 | -34.7867 | -22.8971
13.7346 | -34.7871 | -22.9055
13.7407 | -34.7876 | -22.9139
13.7468 | -34.788 | -22.9224
13.7529 | -34.7884 | -22.9306
13.759 | -34.7889 | -22.9392
13.7651 | -34.7893 | -22.9476
13.7712 | -34.7897 | -22.9562
13.7773 | -34.7902 | -22.9647
13.7834 | -34.7907 | -22.9731
13.7895 | -34.7912 | -22.9814
13.7956 | -34.7916 | -22.99
13.8017 | -34.7921 | -22.9983
13.8078 | -34.7926 | -23.0067
13.8139 | -34.7931 | -23.015
13.82 | -34.7935 | -23.0232
13.8262 | -34.794 | -23.0319
13.8323 | -34.7945 | -23.0404
13.8384 | -34.7949 | -23.049
13.8445 | -34.7954 | -23.0576
13.8506 | -34.7959 | -23.0661
13.8567 | -34.7963 | -23.0748
13.8628 | -34.7967 | -23.0833
13.8689 | -34.7972 | -23.0918
13.875 | -34.7977 | -23.1002
13.8811 | -34.7982 | -23.1085
13.8872 | -34.7986 | -23.1169
13.8933 | -34.7991 | -23.1251
13.8994 | -34.7996 | -23.1333
13.9055 | -34.8 | -23.1416
13.9116 | -34.8005 | -23.1498
13.9177 | -34.801 | -23.1578
13.9238 | -34.8015 | -23.1664
13.9299 | -34.802 | -23.1748
13.936 | -34.8024 | -23.183
13.9421 | -34.8028 | -23.1914
13.9482 | -34.8033 | -23.1998
13.9543 | -34.8038 | -23.2079
13.9604 | -34.8043 | -23.2165
13.9665 | -34.8048 | -23.2249
13.9727 | -34.8052 | -23.2332
13.9788 | -34.8056 | -23.2418
13.9849 | -34.806 | -23.2503
13.991 | -34.8065 | -23.259
13.9971 | -34.8069 | -23.2677
14.0032 | -34.8073 | -23.2763
14.0093 | -34.8078 | -23.2846
14.0154 | -34.8082 | -23.2932
14.0215 | -34.8086 | -23.3014
14.0276 | -34.8091 | -23.3096
14.0337 | -34.8096 | -23.3183
14.0398 | -34.8101 | -23.3264
14.0459 | -34.8105 | -23.3349
14.052 | -34.8109 | -23.3434
14.0581 | -34.8114 | -23.3515
14.0642 | -34.8118 | -23.3601
14.0703 | -34.8123 | -23.3685
14.0764 | -34.8127 | -23.3774
14.0825 | -34.8132 | -23.3858
14.0886 | -34.8137 | -23.3938
14.0947 | -34.8141 | -23.4019
14.1008 | -34.8145 | -23.4106
14.1069 | -34.815 | -23.4189
14.1131 | -34.8154 | -23.4276
14.1192 | -34.8159 | -23.4357
14.1253 | -34.8163 | -23.4438
14.1314 | -34.8168 | -23.4521
14.1375 | -34.8173 | -23.4607
14.1436 | -34.8178 | -23.469
14.1497 | -34.8183 | -23.4771
14.1558 | -34.8187 | -23.4857
14.1619 | -34.8192 | -23.494
14.168 | -34.8196 | -23.5022
14.1741 | -34.8201 | -23.5104
14.1802 | -34.8206 | -23.5187
14.1863 | -34.821 | -23.5271
14.1924 | -34.8215 | -23.5355
14.1985 | -34.822 | -23.5437
14.2046 | -34.8224 | -23.5519
14.2107 | -34.8229 | -23.5598
14.2168 | -34.8234 | -23.5685
14.2229 | -34.8238 | -23.5769
14.229 | -34.8243 | -23.5853
14.2351 | -34.8248 | -23.5934
14.2412 | -34.8252 | -23.6016
14.2473 | -34.8257 | -23.61
14.2534 | -34.8262 | -23.6182
14.2596 | -34.8267 | -23.6266
14.2657 | -34.8271 | -23.6351
14.2718 | -34.8276 | -23.6436
14.2779 | -34.828 | -23.6517
14.284 | -34.8285 | -23.66
14.2901 | -34.829 | -23.6681
14.2962 | -34.8295 | -23.6762
14.3023 | -34.8299 | -23.6846
14.3084 | -34.8304 | -23.6932
14.3145 | -34.8309 | -23.7015
14.3206 | -34.8313 | -23.7099
14.3267 | -34.8318 | -23.7181
14.3328 | -34.8323 | -23.7264
14.3389 | -34.8327 | -23.7345
14.345 | -34.8331 | -23.743
14.3511 | -34.8336 | -23.7511
14.3572 | -34.834 | -23.7595
14.3633 | -34.8345 | -23.7678
14.3694 | -34.8349 | -23.776
14.3755 | -34.8354 | -23.784
14.3816 | -34.8359 | -23.7922
14.3877 | -34.8364 | -23.8003
14.3938 | -34.8368 | -23.8089
14.4 | -34.8373 | -23.8169
14.4061 | -34.8377 | -23.8253
14.4122 | -34.8381 | -23.8338
14.4183 | -34.8385 | -23.8421
14.4244 | -34.839 | -23.8503
14.4305 | -34.8395 | -23.8589
14.4366 | -34.84 | -23.8674
14.4427 | -34.8404 | -23.8759
14.4488 | -34.8409 | -23.8839
14.4549 | -34.8413 | -23.8921
14.461 | -34.8418 | -23.9005
14.4671 | -34.8423 | -23.9089
14.4732 | -34.8428 | -23.9168
14.4793 | -34.8432 | -23.9258
14.4854 | -34.8438 | -23.9337
14.4915 | -34.8442 | -23.9418
14.4976 | -34.8447 | -23.9502
14.5037 | -34.8452 | -23.9586
14.5098 | -34.8457 | -23.9667
14.5159 | -34.8462 | -23.9752
14.522 | -34.8466 | -23.9836
14.5281 | -34.8471 | -23.9916
14.5342 | -34.8475 | -24.0001
14.5403 | -34.848 | -24.0084
14.5465 | -34.8485 | -24.0171
14.5526 | -34.849 | -24.0256
14.5587 | -34.8494 | -24.0339
14.5648 | -34.8499 | -24.042
14.5709 | -34.8504 | -24.0505
14.577 | -34.8508 | -24.0588
14.5831 | -34.8513 | -24.0672
14.5892 | -34.8518 | -24.0751
14.5953 | -34.8523 | -24.0831
14.6014 | -34.8528 | -24.0913
14.6075 | -34.8533 | -24.0993
14.6136 | -34.8537 | -24.1074
14.6197 | -34.8542 | -24.1158
14.6258 | -34.8546 | -24.1242
14.6319 | -34.8552 | -24.1323
14.638 | -34.8557 | -24.1405
14.6441 | -34.8561 | -24.1483
14.6502 | -34.8566 | -24.1564
14.6563 | -34.857 | -24.1646
14.6624 | -34.8575 | -24.1728
14.6685 | -34.858 | -24.1807
14.6746 | -34.8585 | -24.189
14.6807 | -34.8589 | -24.1972
14.6869 | -34.8594 | -24.2053
14.693 | -34.8599 | -24.2134
14.6991 | -34.8603 | -24.2215
14.7052 | -34.8609 | -24.2297
14.7113 | -34.8613 | -24.2382
14.7174 | -34.8618 | -24.2463
14.7235 | -34.8623 | -24.2544
14.7296 | -34.8627 | -24.2625
14.7357 | -34.8632 | -24.2706
14.7418 | -34.8637 | -24.2785
14.7479 | -34.8642 | -24.2867
14.754 | -34.8647 | -24.2949
14.7601 | -34.8651 | -24.3032
14.7662 | -34.8656 | -24.3114
14.7723 | -34.8661 | -24.3194
14.7784 | -34.8665 | -24.3277
14.7845 | -34.867 | -24.336
14.7906 | -34.8675 | -24.3443
14.7967 | -34.868 | -24.3524
14.8028 | -34.8684 | -24.3604
14.8089 | -34.8689 | -24.3685
14.815 | -34.8693 | -24.3765
14.8211 | -34.8698 | -24.3846
14.8272 | -34.8702 | -24.393
14.8334 | -34.8707 | -24.4011
14.8395 | -34.8712 | -24.4092
14.8456 | -34.8717 | -24.4173
14.8517 | -34.8722 | -24.4256
14.8578 | -34.8727 | -24.4337
14.8639 | -34.8732 | -24.4421
14.87 | -34.8737 | -24.4503
14.8761 | -34.8741 | -24.4582
14.8822 | -34.8746 | -24.4664
14.8883 | -34.8751 | -24.4744
14.8944 | -34.8756 | -24.4826
14.9005 | -34.876 | -24.4905
14.9066 | -34.8765 | -24.4987
14.9127 | -34.877 | -24.5067
14.9188 | -34.8775 | -24.5151
14.9249 | -34.878 | -24.5232
14.931 | -34.8784 | -24.5312
14.9371 | -34.8789 | -24.5393
14.9432 | -34.8794 | -24.5476
14.9493 | -34.8799 | -24.5557
14.9554 | -34.8804 | -24.5638
14.9615 | -34.8808 | -24.572
14.9676 | -34.8813 | -24.5803
14.9738 | -34.8818 | -24.5883
14.9799 | -34.8823 | -24.5967
14.986 | -34.8828 | -24.6045
14.9921 | -34.8833 | -24.6124
14.9982 | -34.8838 | -24.6209
15.0043 | -34.8842 | -24.6289
15.0104 | -34.8847 | -24.637
15.0165 | -34.8852 | -24.6453
15.0226 | -34.8857 | -24.653
15.0287 | -34.8861 | -24.6617
15.0348 | -34.8867 | -24.6699
15.0409 | -34.8871 | -24.6783
15.047 | -34.8876 | -24.6867
15.0531 | -34.8882 | -24.6949
15.0592 | -34.8887 | -24.7032
15.0653 | -34.8892 | -24.7112
15.0714 | -34.8897 | -24.7191
15.0775 | -34.8902 | -24.7269
15.0836 | -34.8907 | -24.7351
15.0897 | -34.8912 | -24.7432
15.0958 | -34.8917 | -24.7511
15.1019 | -34.8921 | -24.7592
15.108 | -34.8926 | -24.7674
15.1141 | -34.8931 | -24.7755
15.1203 | -34.8936 | -24.7836
15.1264 | -34.8942 | -24.7919
15.1325 | -34.8947 | -24.7999
15.1386 | -34.8952 | -24.808
15.1447 | -34.8957 | -24.8162
15.1508 | -34.8962 | -24.8246
15.1569 | -34.8967 | -24.833
15.163 | -34.8973 | -24.8406
15.1691 | -34.8978 | -24.8487
15.1752 | -34.8983 | -24.8567
15.1813 | -34.8988 | -24.8649
15.1874 | -34.8993 | -24.8731
15.1935 | -34.8998 | -24.8809
15.1996 | -34.9004 | -24.8892
15.2057 | -34.9008 | -24.8972
15.2118 | -34.9013 | -24.9053
15.2179 | -34.9018 | -24.9132
15.224 | -34.9023 | -24.9213
15.2301 | -34.9028 | -24.9291
15.2362 | -34.9033 | -24.9371
15.2423 | -34.9038 | -24.9455
15.2484 | -34.9043 | -24.9535
15.2545 | -34.9048 | -24.9615
15.2607 | -34.9053 | -24.9696
15.2668 | -34.9059 | -24.9775
15.2729 | -34.9064 | -24.9854
15.279 | -34.9069 | -24.9936
15.2851 | -34.9074 | -25.0017
15.2912 | -34.9079 | -25.0096
15.2973 | -34.9084 | -25.0175
15.3034 | -34.9089 | -25.0254
15.3095 | -34.9094 | -25.0332
15.3156 | -34.9099 | -25.0414
15.3217 | -34.9104 | -25.049
15.3278 | -34.9109 | -25.0572
15.3339 | -34.9114 | -25.065
15.34 | -34.9119 | -25.0728
15.3461 | -34.9124 | -25.081
15.3522 | -34.9129 | -25.0888
15.3583 | -34.9134 | -25.0968
15.3644 | -34.9139 | -25.105
15.3705 | -34.9144 | -25.1127
15.3766 | -34.9149 | -25.1209
15.3827 | -34.9154 | -25.1288
15.3888 | -34.9159 | -25.1365
15.3949 | -34.9164 | -25.1444
15.401 | -34.9169 | -25.1525
15.4072 | -34.9175 | -25.1601
15.4133 | -34.918 | -25.1682
15.4194 | -34.9185 | -25.1762
15.4255 | -34.919 | -25.1843
15.4316 | -34.9195 | -25.1924
15.4377 | -34.92 | -25.2003
15.4438 | -34.9205 | -25.2079
15.4499 | -34.921 | -25.2158
15.456 | -34.9215 | -25.2238
15.4621 | -34.922 | -25.2318
15.4682 | -34.9226 | -25.2399
15.4743 | -34.9231 | -25.2477
15.4804 | -34.9236 | -25.2558
15.4865 | -34.9241 | -25.2637
15.4926 | -34.9246 | -25.2719
15.4987 | -34.9251 | -25.2799
15.5048 | -34.9256 | -25.2878
15.5109 | -34.9262 | -25.2955
15.517 | -34.9267 | -25.3036
15.5231 | -34.9272 | -25.3116
15.5292 | -34.9277 | -25.3194
15.5353 | -34.9282 | -25.3273
15.5414 | -34.9288 | -25.3355
15.5476 | -34.9293 | -25.3436
15.5537 | -34.9298 | -25.3514
15.5598 | -34.9303 | -25.3591
15.5659 | -34.9308 | -25.3673
15.572 | -34.9314 | -25.3748
15.5781 | -34.9318 | -25.3828
15.5842 | -34.9324 | -25.3905
15.5903 | -34.9329 | -25.3981
15.5964 | -34.9335 | -25.4058
15.6025 | -34.934 | -25.4136
15.6086 | -34.9345 | -25.4217
15.6147 | -34.935 | -25.4296
15.6208 | -34.9355 | -25.4372
15.6269 | -34.9361 | -25.445
15.633 | -34.9365 | -25.453
15.6391 | -34.937 | -25.4611
15.6452 | -34.9376 | -25.4689
15.6513 | -34.9381 | -25.477
15.6574 | -34.9386 | -25.4846
15.6635 | -34.9391 | -25.4925
15.6696 | -34.9396 | -25.5004
15.6757 | -34.9402 | -25.5086
15.6818 | -34.9407 | -25.5167
15.688 | -34.9412 | -25.5245
15.6941 | -34.9418 | -25.5325
15.7002 | -34.9423 | -25.5403
15.7063 | -34.9428 | -25.5483
15.7124 | -34.9432 | -25.5563
15.7185 | -34.9438 | -25.5641
15.7246 | -34.9443 | -25.5718
15.7307 | -34.9449 | -25.5798
15.7368 | -34.9454 | -25.588
15.7429 | -34.9459 | -25.5958
15.749 | -34.9464 | -25.6037
15.7551 | -34.9469 | -25.6116
15.7612 | -34.9474 | -25.6189
15.7673 | -34.9479 | -25.6266
15.7734 | -34.9484 | -25.6343
15.7795 | -34.9489 | -25.6425
15.7856 | -34.9495 | -25.6502
15.7917 | -34.95 | -25.6578
15.7978 | -34.9505 | -25.6656
15.8039 | -34.951 | -25.6734
15.81 | -34.9515 | -25.6811
15.8161 | -34.952 | -25.6891
15.8222 | -34.9526 | -25.6969
15.8283 | -34.9531 | -25.7046
15.8345 | -34.9536 | -25.7125
15.8406 | -34.9542 | -25.7203
15.8467 | -34.9547 | -25.7282
15.8528 | -34.9553 | -25.736
15.8589 | -34.9558 | -25.7437
15.865 | -34.9563 | -25.7515
15.8711 | -34.9568 | -25.7596
15.8772 | -34.9574 | -25.7675
15.8833 | -34.9579 | -25.7757
15.8894 | -34.9584 | -25.7834
15.8955 | -34.959 | -25.7913
15.9016 | -34.9595 | -25.7991
15.9077 | -34.9601 | -25.8069
15.9138 | -34.9606 | -25.8146
15.9199 | -34.9611 | -25.8226
15.926 | -34.9616 | -25.8307
15.9321 | -34.9622 | -25.8385
15.9382 | -34.9628 | -25.8462
15.9443 | -34.9633 | -25.8542
15.9504 | -34.9638 | -25.862
15.9565 | -34.9644 | -25.8698
15.9626 | -34.965 | -25.8776
15.9687 | -34.9655 | -25.8852
15.9749 | -34.9661 | -25.8931
15.981 | -34.9666 | -25.901
15.9871 | -34.9671 | -25.9088
15.9932 | -34.9676 | -25.9167
15.9993 | -34.9682 | -25.9243
16.0054 | -34.9687 | -25.9318
16.0115 | -34.9693 | -25.9395
16.0176 | -34.9698 | -25.9476
16.0237 | -34.9703 | -25.9557
16.0298 | -34.9709 | -25.9631
16.0359 | -34.9714 | -25.9707
16.042 | -34.9719 | -25.9785
16.0481 | -34.9724 | -25.9863
16.0542 | -34.973 | -25.9944
16.0603 | -34.9735 | -26.002
16.0664 | -34.9741 | -26.0097
16.0725 | -34.9746 | -26.0173
16.0786 | -34.9752 | -26.0253
16.0847 | -34.9757 | -26.0328
16.0908 | -34.9762 | -26.0405
16.0969 | -34.9768 | -26.0481
16.103 | -34.9772 | -26.0561
16.1091 | -34.9778 | -26.0637
16.1152 | -34.9783 | -26.0716
16.1214 | -34.9789 | -26.0794
16.1275 | -34.9794 | -26.087
16.1336 | -34.9799 | -26.0948
16.1397 | -34.9805 | -26.1025
16.1458 | -34.981 | -26.1102
16.1519 | -34.9816 | -26.1181
16.158 | -34.9821 | -26.126
16.1641 | -34.9826 | -26.1338
16.1702 | -34.9831 | -26.1417
16.1763 | -34.9837 | -26.1495
16.1824 | -34.9842 | -26.157
16.1885 | -34.9848 | -26.1647
16.1946 | -34.9853 | -26.1724
16.2007 | -34.9858 | -26.1805
16.2068 | -34.9864 | -26.1881
16.2129 | -34.9869 | -26.196
16.219 | -34.9874 | -26.2036
16.2251 | -34.9879 | -26.2115
16.2312 | -34.9884 | -26.2193
16.2373 | -34.9889 | -26.227
16.2434 | -34.9894 | -26.2347
16.2495 | -34.99 | -26.2422
16.2556 | -34.9906 | -26.25
16.2618 | -34.9911 | -26.2576
16.2679 | -34.9917 | -26.2655
16.274 | -34.9922 | -26.2734
16.2801 | -34.9928 | -26.2808
16.2862 | -34.9933 | -26.2884
16.2923 | -34.9938 | -26.2963
16.2984 | -34.9943 | -26.3042
16.3045 | -34.9949 | -26.3118
16.3106 | -34.9955 | -26.3192
16.3167 | -34.996 | -26.3271
16.3228 | -34.9965 | -26.3348
16.3289 | -34.997 | -26.3425
16.335 | -34.9976 | -26.3504
16.3411 | -34.9981 | -26.3579
16.3472 | -34.9986 | -26.3658
16.3533 | -34.9992 | -26.3736
16.3594 | -34.9997 | -26.3811
16.3655 | -35.0002 | -26.389
16.3716 | -35.0007 | -26.3966
16.3777 | -35.0013 | -26.4043
16.3838 | -35.0018 | -26.4119
16.3899 | -35.0023 | -26.4197
16.396 | -35.0028 | -26.4277
16.4021 | -35.0034 | -26.4354
16.4083 | -35.004 | -26.443
16.4144 | -35.0045 | -26.4505
16.4205 | -35.005 | -26.4584
16.4266 | -35.0056 | -26.4662
16.4327 | -35.0061 | -26.4737
16.4388 | -35.0066 | -26.4815
16.4449 | -35.0072 | -26.4891
16.451 | -35.0077 | -26.497
16.4571 | -35.0082 | -26.5049
16.4632 | -35.0087 | -26.5124
16.4693 | -35.0092 | -26.5202
16.4754 | -35.0098 | -26.528
16.4815 | -35.0103 | -26.5359
16.4876 | -35.0108 | -26.5435
16.4937 | -35.0114 | -26.5511
16.4998 | -35.012 | -26.5588
16.5059 | -35.0125 | -26.5664
16.512 | -35.0131 | -26.5745
16.5181 | -35.0136 | -26.5823
16.5242 | -35.0141 | -26.5901
16.5303 | -35.0147 | -26.5976
16.5364 | -35.0152 | -26.6052
16.5425 | -35.0157 | -26.6131
16.5487 | -35.0162 | -26.6209
16.5548 | -35.0168 | -26.6285
16.5609 | -35.0173 | -26.6364
16.567 | -35.0178 | -26.6441
16.5731 | -35.0184 | -26.6518
16.5792 | -35.0189 | -26.6596
16.5853 | -35.0195 | -26.6672
16.5914 | -35.02 | -26.675
16.5975 | -35.0205 | -26.6827
16.6036 | -35.0211 | -26.6903
16.6097 | -35.0216 | -26.6979
16.6158 | -35.0221 | -26.7058
16.6219 | -35.0227 | -26.7134
16.628 | -35.0232 | -26.7212
16.6341 | -35.0237 | -26.7287
16.6402 | -35.0243 | -26.7359
16.6463 | -35.0248 | -26.7436
16.6524 | -35.0253 | -26.7512
16.6585 | -35.0259 | -26.7591
16.6646 | -35.0264 | -26.767
16.6707 | -35.0269 | -26.7744
16.6768 | -35.0275 | -26.7819
16.6829 | -35.028 | -26.7895
16.689 | -35.0285 | -26.7973
16.6952 | -35.029 | -26.8051
16.7013 | -35.0296 | -26.8127
16.7074 | -35.0301 | -26.8202
16.7135 | -35.0306 | -26.8278
16.7196 | -35.0311 | -26.8355
16.7257 | -35.0316 | -26.8431
16.7318 | -35.0322 | -26.8508
16.7379 | -35.0327 | -26.8585
16.744 | -35.0333 | -26.8659
16.7501 | -35.0338 | -26.8737
16.7562 | -35.0343 | -26.8815
16.7623 | -35.0349 | -26.8891
16.7684 | -35.0354 | -26.897
16.7745 | -35.036 | -26.9044
16.7806 | -35.0365 | -26.912
16.7867 | -35.037 | -26.9196
16.7928 | -35.0375 | -26.9272
16.7989 | -35.038 | -26.9347
16.805 | -35.0385 | -26.9424
16.8111 | -35.0391 | -26.9501
16.8172 | -35.0396 | -26.9582
16.8233 | -35.0402 | -26.9662
16.8294 | -35.0407 | -26.9739
16.8356 | -35.0412 | -26.9817
16.8417 | -35.0418 | -26.9894
16.8478 | -35.0423 | -26.9971
16.8539 | -35.0428 | -27.0047
16.86 | -35.0434 | -27.0122
16.8661 | -35.0439 | -27.02
16.8722 | -35.0445 | -27.0276
16.8783 | -35.045 | -27.035
16.8844 | -35.0455 | -27.0426
16.8905 | -35.0461 | -27.0504
16.8966 | -35.0466 | -27.0581
16.9027 | -35.0471 | -27.0656
16.9088 | -35.0477 | -27.0736
16.9149 | -35.0482 | -27.0816
16.921 | -35.0487 | -27.0893
16.9271 | -35.0493 | -27.0972
16.9332 | -35.0498 | -27.1049
16.9393 | -35.0503 | -27.1129
16.9454 | -35.0508 | -27.1204
16.9515 | -35.0513 | -27.1283
16.9576 | -35.0519 | -27.1358
16.9637 | -35.0525 | -27.1435
16.9698 | -35.053 | -27.1509
16.9759 | -35.0535 | -27.1582
16.9821 | -35.0541 | -27.166
16.9882 | -35.0546 | -27.1739
16.9943 | -35.0551 | -27.1816
17.0004 | -35.0557 | -27.1893
17.0065 | -35.0562 | -27.197
17.0126 | -35.0567 | -27.2045
17.0187 | -35.0573 | -27.2122
17.0248 | -35.0578 | -27.2196
17.0309 | -35.0583 | -27.2273
17.037 | -35.0588 | -27.235
17.0431 | -35.0594 | -27.2424
17.0492 | -35.0599 | -27.2501
17.0553 | -35.0604 | -27.2576
17.0614 | -35.0609 | -27.2652
17.0675 | -35.0614 | -27.2731
17.0736 | -35.062 | -27.2809
17.0797 | -35.0625 | -27.2881
17.0858 | -35.063 | -27.2958
17.0919 | -35.0635 | -27.3036
17.098 | -35.064 | -27.3109
17.1041 | -35.0646 | -27.3183
17.1102 | -35.0651 | -27.3262
17.1163 | -35.0656 | -27.334
17.1225 | -35.0662 | -27.3413
17.1286 | -35.0667 | -27.3487
17.1347 | -35.0672 | -27.3563
17.1408 | -35.0678 | -27.364
17.1469 | -35.0683 | -27.3715
17.153 | -35.0689 | -27.3789
17.1591 | -35.0694 | -27.3862
17.1652 | -35.0699 | -27.3937
17.1713 | -35.0704 | -27.4013
17.1774 | -35.071 | -27.4091
17.1835 | -35.0715 | -27.4168
17.1896 | -35.072 | -27.4242
17.1957 | -35.0725 | -27.432
17.2018 | -35.0731 | -27.4395
17.2079 | -35.0736 | -27.4471
17.214 | -35.0741 | -27.4547
17.2201 | -35.0746 | -27.4623
17.2262 | -35.0752 | -27.4699
17.2323 | -35.0757 | -27.4775
17.2384 | -35.0762 | -27.4851
17.2445 | -35.0767 | -27.4925
17.2506 | -35.0772 | -27.5003
17.2567 | -35.0777 | -27.5079
17.2628 | -35.0783 | -27.5158
17.269 | -35.0788 | -27.5233
17.2751 | -35.0793 | -27.531
17.2812 | -35.0798 | -27.5385
17.2873 | -35.0804 | -27.5461
17.2934 | -35.0808 | -27.5536
17.2995 | -35.0813 | -27.5611
17.3056 | -35.0818 | -27.5688
17.3117 | -35.0823 | -27.5763
17.3178 | -35.0829 | -27.5838
17.3239 | -35.0833 | -27.5914
17.33 | -35.0839 | -27.599
17.3361 | -35.0843 | -27.6066
17.3422 | -35.0849 | -27.6142
17.3483 | -35.0854 | -27.6219
17.3544 | -35.0859 | -27.6295
17.3605 | -35.0865 | -27.6371
17.3666 | -35.087 | -27.6449
17.3727 | -35.0875 | -27.6524
17.3788 | -35.088 | -27.6597
17.3849 | -35.0886 | -27.6672
17.391 | -35.0891 | -27.6744
17.3971 | -35.0896 | -27.6821
17.4032 | -35.0901 | -27.6901
17.4094 | -35.0907 | -27.6974
17.4155 | -35.0912 | -27.7049
17.4216 | -35.0917 | -27.7124
17.4277 | -35.0922 | -27.7202
17.4338 | -35.0927 | -27.7277
17.4399 | -35.0932 | -27.7353
17.446 | -35.0937 | -27.7428
17.4521 | -35.0942 | -27.7503
17.4582 | -35.0947 | -27.7577
17.4643 | -35.0953 | -27.7654
17.4704 | -35.0958 | -27.7727
17.4765 | -35.0963 | -27.7803
17.4826 | -35.0968 | -27.7878
17.4887 | -35.0973 | -27.7956
17.4948 | -35.0979 | -27.8031
17.5009 | -35.0983 | -27.8104
17.507 | -35.0989 | -27.8179
17.5131 | -35.0994 | -27.8254
17.5192 | -35.0998 | -27.833
17.5253 | -35.1003 | -27.8406
17.5314 | -35.1008 | -27.8481
17.5375 | -35.1013 | -27.8558
17.5436 | -35.1018 | -27.8634
17.5497 | -35.1023 | -27.8712
17.5559 | -35.1028 | -27.8788
17.562 | -35.1033 | -27.8862
17.5681 | -35.1039 | -27.8936
17.5742 | -35.1043 | -27.9012
17.5803 | -35.1049 | -27.9084
17.5864 | -35.1054 | -27.9158
17.5925 | -35.106 | -27.9237
17.5986 | -35.1065 | -27.931
17.6047 | -35.107 | -27.9388
17.6108 | -35.1076 | -27.946
17.6169 | -35.1081 | -27.9535
17.623 | -35.1085 | -27.9609
17.6291 | -35.1091 | -27.9683
17.6352 | -35.1096 | -27.9758
17.6413 | -35.1101 | -27.9836
17.6474 | -35.1106 | -27.9913
17.6535 | -35.1111 | -27.9989
17.6596 | -35.1117 | -28.0063
17.6657 | -35.1122 | -28.0137
17.6718 | -35.1127 | -28.0214
17.6779 | -35.1132 | -28.0288
17.684 | -35.1138 | -28.0364
17.6901 | -35.1143 | -28.0438
17.6963 | -35.1148 | -28.051
17.7024 | -35.1153 | -28.0584
17.7085 | -35.1158 | -28.0657
17.7146 | -35.1164 | -28.0733
17.7207 | -35.1169 | -28.0806
17.7268 | -35.1174 | -28.088
17.7329 | -35.1179 | -28.0959
17.739 | -35.1184 | -28.1035
17.7451 | -35.119 | -28.1109
17.7512 | -35.1195 | -28.1186
17.7573 | -35.12 | -28.1261
17.7634 | -35.1205 | -28.1334
17.7695 | -35.121 | -28.141
17.7756 | -35.1215 | -28.1485
17.7817 | -35.122 | -28.1557
17.7878 | -35.1225 | -28.1633
17.7939 | -35.1231 | -28.1707
17.8 | -35.1236 | -28.1781
17.8061 | -35.1241 | -28.1857
17.8122 | -35.1246 | -28.1933
17.8183 | -35.1251 | -28.201
17.8244 | -35.1256 | -28.2084
17.8305 | -35.1262 | -28.2159
17.8366 | -35.1267 | -28.2231
17.8428 | -35.1272 | -28.2309
17.8489 | -35.1277 | -28.2382
17.855 | -35.1283 | -28.2455
17.8611 | -35.1288 | -28.2529
17.8672 | -35.1293 | -28.2602
17.8733 | -35.1298 | -28.268
17.8794 | -35.1303 | -28.2756
17.8855 | -35.1309 | -28.2832
17.8916 | -35.1314 | -28.2904
17.8977 | -35.1319 | -28.2979
17.9038 | -35.1324 | -28.3051
17.9099 | -35.1329 | -28.3124
17.916 | -35.1334 | -28.3199
17.9221 | -35.1339 | -28.3273
17.9282 | -35.1344 | -28.3347
17.9343 | -35.1349 | -28.3422
17.9404 | -35.1354 | -28.3494
17.9465 | -35.1359 | -28.3569
17.9526 | -35.1364 | -28.3646
17.9587 | -35.137 | -28.3721
17.9648 | -35.1375 | -28.3795
17.9709 | -35.138 | -28.3869
17.977 | -35.1385 | -28.3941
17.9832 | -35.139 | -28.4016
17.9893 | -35.1395 | -28.4091
17.9954 | -35.14 | -28.4167
18.0015 | -35.1406 | -28.4242
18.0076 | -35.141 | -28.4312
18.0137 | -35.1415 | -28.4389
18.0198 | -35.142 | -28.4463
18.0259 | -35.1425 | -28.4539
18.032 | -35.1429 | -28.4612
18.0381 | -35.1435 | -28.4685
18.0442 | -35.144 | -28.4761
18.0503 | -35.1445 | -28.4836
18.0564 | -35.1449 | -28.4911
18.0625 | -35.1454 | -28.4987
18.0686 | -35.146 | -28.5061
18.0747 | -35.1465 | -28.5135
18.0808 | -35.147 | -28.5208
18.0869 | -35.1475 | -28.5282
18.093 | -35.148 | -28.5356
18.0991 | -35.1485 | -28.5434
18.1052 | -35.1491 | -28.551
18.1113 | -35.1496 | -28.5581
18.1174 | -35.1501 | -28.5655
18.1236 | -35.1507 | -28.573
18.1297 | -35.1512 | -28.5804
18.1358 | -35.1517 | -28.5878
18.1419 | -35.1523 | -28.5955
18.148 | -35.1528 | -28.6031
18.1541 | -35.1533 | -28.6102
18.1602 | -35.1539 | -28.6178
18.1663 | -35.1544 | -28.6253
18.1724 | -35.1549 | -28.633
18.1785 | -35.1554 | -28.6405
18.1846 | -35.1559 | -28.6478
18.1907 | -35.1564 | -28.6556
18.1968 | -35.1568 | -28.6631
18.2029 | -35.1573 | -28.6701
18.209 | -35.1579 | -28.6778
18.2151 | -35.1584 | -28.6853
18.2212 | -35.1589 | -28.6924
18.2273 | -35.1595 | -28.6997
18.2334 | -35.16 | -28.7071
18.2395 | -35.1605 | -28.7144
18.2456 | -35.161 | -28.7218
18.2517 | -35.1615 | -28.7292
18.2578 | -35.162 | -28.7365
18.2639 | -35.1625 | -28.7441
18.2701 | -35.1631 | -28.7515
18.2762 | -35.1636 | -28.7591
18.2823 | -35.1641 | -28.7665
18.2884 | -35.1646 | -28.774
18.2945 | -35.1652 | -28.7812
18.3006 | -35.1657 | -28.7887
18.3067 | -35.1662 | -28.796
18.3128 | -35.1667 | -28.8034
18.3189 | -35.1672 | -28.8111
18.325 | -35.1677 | -28.818
18.3311 | -35.1682 | -28.8256
18.3372 | -35.1687 | -28.8327
18.3433 | -35.1692 | -28.84
18.3494 | -35.1697 | -28.8475
18.3555 | -35.1702 | -28.8547
18.3616 | -35.1707 | -28.8617
18.3677 | -35.1713 | -28.8689
18.3738 | -35.1718 | -28.8761
18.3799 | -35.1723 | -28.8834
18.386 | -35.1728 | -28.8907
18.3921 | -35.1733 | -28.8979
18.3982 | -35.1738 | -28.9051
18.4043 | -35.1742 | -28.9123
18.4105 | -35.1747 | -28.9197
18.4166 | -35.1752 | -28.9269
18.4227 | -35.1758 | -28.9344
18.4288 | -35.1763 | -28.942
18.4349 | -35.1768 | -28.9492
18.441 | -35.1773 | -28.9565
18.4471 | -35.1778 | -28.9639
18.4532 | -35.1784 | -28.9716
18.4593 | -35.1789 | -28.9788
18.4654 | -35.1794 | -28.9857
18.4715 | -35.1799 | -28.9931
18.4776 | -35.1804 | -29.0006
18.4837 | -35.181 | -29.0081
18.4898 | -35.1815 | -29.0157
18.4959 | -35.182 | -29.0231
18.502 | -35.1825 | -29.0305
18.5081 | -35.1831 | -29.0381
18.5142 | -35.1835 | -29.0456
18.5203 | -35.1841 | -29.0527
18.5264 | -35.1846 | -29.0599
18.5325 | -35.1851 | -29.0675
18.5386 | -35.1857 | -29.075
18.5447 | -35.1862 | -29.0823
18.5508 | -35.1867 | -29.0895
18.557 | -35.1872 | -29.097
18.5631 | -35.1877 | -29.1043
18.5692 | -35.1882 | -29.1114
18.5753 | -35.1888 | -29.1186
18.5814 | -35.1893 | -29.1262
18.5875 | -35.1898 | -29.1338
18.5936 | -35.1903 | -29.1416
18.5997 | -35.1908 | -29.1489
18.6058 | -35.1913 | -29.1562
18.6119 | -35.1919 | -29.1637
18.618 | -35.1924 | -29.1711
18.6241 | -35.193 | -29.1787
18.6302 | -35.1935 | -29.1864
18.6363 | -35.194 | -29.1934
18.6424 | -35.1945 | -29.2007
18.6485 | -35.195 | -29.2079
18.6546 | -35.1955 | -29.215
18.6607 | -35.196 | -29.2222
18.6668 | -35.1966 | -29.2295
18.6729 | -35.1971 | -29.237
18.679 | -35.1976 | -29.2441
18.6851 | -35.1981 | -29.2513
18.6912 | -35.1986 | -29.2585
18.6974 | -35.1991 | -29.2659
18.7035 | -35.1996 | -29.2731
18.7096 | -35.2001 | -29.2805
18.7157 | -35.2007 | -29.2877
18.7218 | -35.2012 | -29.295
18.7279 | -35.2017 | -29.3021
18.734 | -35.2023 | -29.3096
18.7401 | -35.2028 | -29.3166
18.7462 | -35.2033 | -29.3235
18.7523 | -35.2039 | -29.331
18.7584 | -35.2044 | -29.3384
18.7645 | -35.205 | -29.3456
18.7706 | -35.2054 | -29.3529
18.7767 | -35.2059 | -29.36
18.7828 | -35.2065 | -29.3667
18.7889 | -35.2069 | -29.374
18.795 | -35.2074 | -29.3811
18.8011 | -35.2079 | -29.3883
18.8072 | -35.2085 | -29.3956
18.8133 | -35.209 | -29.4029
18.8194 | -35.2095 | -29.4101
18.8255 | -35.21 | -29.4172
18.8316 | -35.2105 | -29.4246
18.8377 | -35.211 | -29.4318
18.8439 | -35.2115 | -29.4391
18.85 | -35.212 | -29.4464
18.8561 | -35.2126 | -29.4535
18.8622 | -35.213 | -29.461
18.8683 | -35.2136 | -29.4681
18.8744 | -35.2141 | -29.4753
18.8805 | -35.2146 | -29.4827
18.8866 | -35.2151 | -29.4902
18.8927 | -35.2157 | -29.4973
18.8988 | -35.2162 | -29.5044
18.9049 | -35.2167 | -29.5116
18.911 | -35.2173 | -29.5186
18.9171 | -35.2178 | -29.5259
18.9232 | -35.2184 | -29.5328
18.9293 | -35.2188 | -29.5404
18.9354 | -35.2193 | -29.5473
18.9415 | -35.2198 | -29.5544
18.9476 | -35.2203 | -29.5618
18.9537 | -35.2209 | -29.5691
18.9598 | -35.2214 | -29.5757
18.9659 | -35.222 | -29.5828
18.972 | -35.2225 | -29.5901
18.9781 | -35.2229 | -29.5972
18.9843 | -35.2235 | -29.6046
18.9904 | -35.224 | -29.6119
18.9965 | -35.2245 | -29.6191
19.0026 | -35.2251 | -29.6263
19.0087 | -35.2256 | -29.6337
19.0148 | -35.2261 | -29.6408
19.0209 | -35.2266 | -29.6484
19.027 | -35.2272 | -29.6557
19.0331 | -35.2277 | -29.663
19.0392 | -35.2283 | -29.6703
19.0453 | -35.2288 | -29.6777
19.0514 | -35.2293 | -29.6849
19.0575 | -35.2298 | -29.6922
19.0636 | -35.2303 | -29.6987
19.0697 | -35.231 | -29.7064
19.0758 | -35.2315 | -29.7134
19.0819 | -35.232 | -29.7209
19.088 | -35.2325 | -29.7276
19.0941 | -35.233 | -29.7344
19.1002 | -35.2335 | -29.7416
19.1063 | -35.2341 | -29.7489
19.1124 | -35.2347 | -29.7562
19.1185 | -35.2351 | -29.7635
19.1246 | -35.2356 | -29.7704
19.1308 | -35.2361 | -29.7774
19.1369 | -35.2367 | -29.7841
19.143 | -35.2372 | -29.7914
19.1491 | -35.2377 | -29.7982
19.1552 | -35.2382 | -29.8057
19.1613 | -35.2387 | -29.8128
19.1674 | -35.2392 | -29.8201
19.1735 | -35.2398 | -29.8272
19.1796 | -35.2403 | -29.8342
19.1857 | -35.2408 | -29.8413
19.1918 | -35.2413 | -29.8486
19.1979 | -35.2418 | -29.856
19.204 | -35.2423 | -29.8628
19.2101 | -35.2429 | -29.8701
19.2162 | -35.2433 | -29.8776
19.2223 | -35.2438 | -29.8844
19.2284 | -35.2443 | -29.8917
19.2345 | -35.2449 | -29.899
19.2406 | -35.2454 | -29.9062
19.2467 | -35.2459 | -29.9133
19.2528 | -35.2464 | -29.9205
19.2589 | -35.2469 | -29.9277
19.265 | -35.2474 | -29.9347
19.2712 | -35.2479 | -29.9419
19.2773 | -35.2485 | -29.9487
19.2834 | -35.249 | -29.9557
19.2895 | -35.2495 | -29.9629
19.2956 | -35.2501 | -29.9702
19.3017 | -35.2506 | -29.977
19.3078 | -35.2511 | -29.9842
19.3139 | -35.2516 | -29.9911
19.32 | -35.2522 | -29.9982
19.3261 | -35.2526 | -30.0053
19.3322 | -35.2532 | -30.0124
19.3383 | -35.2538 | -30.0192
19.3444 | -35.2543 | -30.0266
19.3505 | -35.2548 | -30.0337
19.3566 | -35.2552 | -30.0407
19.3627 | -35.2558 | -30.0479
19.3688 | -35.2563 | -30.0551
19.3749 | -35.2568 | -30.0624
19.381 | -35.2573 | -30.0693
19.3871 | -35.2578 | -30.0765
19.3932 | -35.2584 | -30.0839
19.3993 | -35.2589 | -30.091
19.4054 | -35.2594 | -30.0982
19.4115 | -35.2599 | -30.1054
19.4177 | -35.2605 | -30.1122
19.4238 | -35.261 | -30.1193
19.4299 | -35.2615 | -30.1261
19.436 | -35.2621 | -30.1335
19.4421 | -35.2626 | -30.14
19.4482 | -35.2632 | -30.1472
19.4543 | -35.2637 | -30.1541
19.4604 | -35.2642 | -30.1612
19.4665 | -35.2648 | -30.1683
19.4726 | -35.2653 | -30.175
19.4787 | -35.2658 | -30.1821
19.4848 | -35.2663 | -30.189
19.4909 | -35.2668 | -30.1958
19.497 | -35.2673 | -30.203
19.5031 | -35.2678 | -30.2098
19.5092 | -35.2683 | -30.2168
19.5153 | -35.2689 | -30.2239
19.5214 | -35.2694 | -30.2309
19.5275 | -35.2699 | -30.238
19.5336 | -35.2704 | -30.2448
19.5397 | -35.271 | -30.2518
19.5458 | -35.2715 | -30.2588
19.5519 | -35.272 | -30.2658
19.5581 | -35.2725 | -30.2728
19.5642 | -35.2731 | -30.28
19.5703 | -35.2736 | -30.287
19.5764 | -35.2741 | -30.294
19.5825 | -35.2746 | -30.3009
19.5886 | -35.2751 | -30.3079
19.5947 | -35.2756 | -30.315
19.6008 | -35.2761 | -30.3222
19.6069 | -35.2766 | -30.3291
19.613 | -35.2771 | -30.3364
19.6191 | -35.2777 | -30.3432
19.6252 | -35.2782 | -30.3501
19.6313 | -35.2787 | -30.3567
19.6374 | -35.2792 | -30.3638
19.6435 | -35.2797 | -30.3707
19.6496 | -35.2802 | -30.3776
19.6557 | -35.2807 | -30.3848
19.6618 | -35.2812 | -30.3916
19.6679 | -35.2817 | -30.3988
19.674 | -35.2822 | -30.4059
19.6801 | -35.2827 | -30.4131
19.6862 | -35.2833 | -30.4199
19.6923 | -35.2838 | -30.4268
19.6984 | -35.2844 | -30.4337
19.7046 | -35.285 | -30.4407
19.7107 | -35.2855 | -30.4477
19.7168 | -35.286 | -30.4548
19.7229 | -35.2865 | -30.4617
19.729 | -35.287 | -30.4686
19.7351 | -35.2875 | -30.4752
19.7412 | -35.2881 | -30.4822
19.7473 | -35.2886 | -30.4894
19.7534 | -35.2891 | -30.4965
19.7595 | -35.2896 | -30.5036
19.7656 | -35.2901 | -30.5108
19.7717 | -35.2906 | -30.5173
19.7778 | -35.2911 | -30.5245
19.7839 | -35.2916 | -30.5315
19.79 | -35.2922 | -30.5383
19.7961 | -35.2926 | -30.5454
19.8022 | -35.2931 | -30.5524
19.8083 | -35.2936 | -30.5593
19.8144 | -35.2942 | -30.5662
19.8205 | -35.2947 | -30.5727
19.8266 | -35.2953 | -30.5799
19.8327 | -35.2957 | -30.5871
19.8388 | -35.2962 | -30.5941
19.845 | -35.2967 | -30.6012
19.8511 | -35.2972 | -30.6083
19.8572 | -35.2978 | -30.6155
19.8633 | -35.2983 | -30.6224
19.8694 | -35.2989 | -30.6299
19.8755 | -35.2994 | -30.637
19.8816 | -35.2999 | -30.6441
19.8877 | -35.3004 | -30.6511
19.8938 | -35.3009 | -30.658
19.8999 | -35.3015 | -30.6649
19.906 | -35.302 | -30.672
19.9121 | -35.3025 | -30.6787
19.9182 | -35.303 | -30.6852
19.9243 | -35.3036 | -30.6923
19.9304 | -35.304 | -30.6992
19.9365 | -35.3045 | -30.7061
19.9426 | -35.305 | -30.7128
19.9487 | -35.3056 | -30.7196
19.9548 | -35.3061 | -30.7268
19.9609 | -35.3067 | -30.7339
19.967 | -35.3072 | -30.7406
19.9731 | -35.3077 | -30.7473
19.9792 | -35.3082 | -30.7538
19.9853 | -35.3087 | -30.7605
19.9915 | -35.3093 | -30.7676
19.9976 | -35.3098 | -30.7741
20.0037 | -35.3103 | -30.7808
20.0098 | -35.3108 | -30.7877
20.0159 | -35.3114 | -30.7946
20.022 | -35.3119 | -30.802
20.0281 | -35.3124 | -30.8091
20.0342 | -35.313 | -30.816
20.0403 | -35.3135 | -30.8224
20.0464 | -35.314 | -30.8298
20.0525 | -35.3146 | -30.837
20.0586 | -35.3151 | -30.8437
20.0647 | -35.3156 | -30.851
20.0708 | -35.3161 | -30.8579
20.0769 | -35.3166 | -30.8648
20.083 | -35.3171 | -30.8717
20.0891 | -35.3176 | -30.8787
20.0952 | -35.3182 | -30.8855
20.1013 | -35.3186 | -30.8923
20.1074 | -35.319 | -30.8995
20.1135 | -35.3196 | -30.9064
20.1196 | -35.3201 | -30.9133
20.1257 | -35.3207 | -30.9202
20.1319 | -35.3212 | -30.9272
20.138 | -35.3217 | -30.9341
20.1441 | -35.3221 | -30.9413
20.1502 | -35.3227 | -30.9485
20.1563 | -35.3232 | -30.9549
20.1624 | -35.3237 | -30.9615
20.1685 | -35.3242 | -30.9683
20.1746 | -35.3247 | -30.9755
20.1807 | -35.3252 | -30.9825
20.1868 | -35.3257 | -30.9895
20.1929 | -35.3262 | -30.9962
20.199 | -35.3267 | -31.0033
20.2051 | -35.3272 | -31.0099
20.2112 | -35.3278 | -31.0167
20.2173 | -35.3282 | -31.0239
20.2234 | -35.3287 | -31.0308
20.2295 | -35.3293 | -31.0377
20.2356 | -35.3298 | -31.0446
20.2417 | -35.3303 | -31.0513
20.2478 | -35.3307 | -31.0583
20.2539 | -35.3312 | -31.0653
20.26 | -35.3317 | -31.0719
20.2661 | -35.3322 | -31.0788
20.2723 | -35.3327 | -31.0857
20.2784 | -35.3332 | -31.0926
20.2845 | -35.3338 | -31.0992
20.2906 | -35.3343 | -31.1059
20.2967 | -35.3348 | -31.1122
20.3028 | -35.3353 | -31.119
20.3089 | -35.3357 | -31.126
20.315 | -35.3363 | -31.1329
20.3211 | -35.3368 | -31.1399
20.3272 | -35.3373 | -31.1466
20.3333 | -35.3379 | -31.1532
20.3394 | -35.3385 | -31.1598
20.3455 | -35.339 | -31.1664
20.3516 | -35.3395 | -31.1731
20.3577 | -35.34 | -31.1796
20.3638 | -35.3405 | -31.1864
20.3699 | -35.3411 | -31.1929
20.376 | -35.3416 | -31.1998
20.3821 | -35.3421 | -31.2065
20.3882 | -35.3426 | -31.2135
20.3943 | -35.3431 | -31.2205
20.4004 | -35.3436 | -31.2271
20.4065 | -35.3441 | -31.2337
20.4126 | -35.3446 | -31.24
20.4188 | -35.3451 | -31.247
20.4249 | -35.3457 | -31.254
20.431 | -35.3461 | -31.2606
20.4371 | -35.3467 | -31.2669
20.4432 | -35.3472 | -31.2738
20.4493 | -35.3477 | -31.2806
20.4554 | -35.3482 | -31.2873
20.4615 | -35.3487 | -31.2942
20.4676 | -35.3492 | -31.3008
20.4737 | -35.3496 | -31.3073
20.4798 | -35.3502 | -31.314
20.4859 | -35.3507 | -31.3205
20.492 | -35.3512 | -31.3273
20.4981 | -35.3517 | -31.3338
20.5042 | -35.3522 | -31.3402
20.5103 | -35.3528 | -31.3467
20.5164 | -35.3533 | -31.3534
20.5225 | -35.3537 | -31.3605
20.5286 | -35.3542 | -31.3676
20.5347 | -35.3547 | -31.374
20.5408 | -35.3552 | -31.3809
20.5469 | -35.3557 | -31.3876
20.553 | -35.3562 | -31.3944
20.5592 | -35.3567 | -31.4011
20.5653 | -35.3572 | -31.4081
20.5714 | -35.3577 | -31.4143
20.5775 | -35.3582 | -31.4209
20.5836 | -35.3587 | -31.4275
20.5897 | -35.3592 | -31.4342
20.5958 | -35.3597 | -31.4409
20.6019 | -35.3602 | -31.4476
20.608 | -35.3607 | -31.4546
20.6141 | -35.3612 | -31.4614
20.6202 | -35.3617 | -31.4682
20.6263 | -35.3622 | -31.4748
20.6324 | -35.3627 | -31.4815
20.6385 | -35.3632 | -31.488
20.6446 | -35.3636 | -31.495
20.6507 | -35.3641 | -31.5014
20.6568 | -35.3647 | -31.5079
20.6629 | -35.3652 | -31.5142
20.669 | -35.3657 | -31.5206
20.6751 | -35.3662 | -31.5269
20.6812 | -35.3668 | -31.5331
20.6873 | -35.3673 | -31.5397
20.6934 | -35.3678 | -31.5465
20.6995 | -35.3683 | -31.553
20.7057 | -35.3688 | -31.5596
20.7118 | -35.3693 | -31.5662
20.7179 | -35.3698 | -31.5726
20.724 | -35.3703 | -31.579
20.7301 | -35.3707 | -31.5856
20.7362 | -35.3713 | -31.5924
20.7423 | -35.3718 | -31.599
20.7484 | -35.3723 | -31.6059
20.7545 | -35.3728 | -31.6124
20.7606 | -35.3733 | -31.6193
20.7667 | -35.3739 | -31.6262
20.7728 | -35.3744 | -31.6329
20.7789 | -35.3749 | -31.6393
20.785 | -35.3754 | -31.646
20.7911 | -35.3759 | -31.6523
20.7972 | -35.3764 | -31.6588
20.8033 | -35.3769 | -31.6652
20.8094 | -35.3774 | -31.672
20.8155 | -35.3779 | -31.6787
20.8216 | -35.3784 | -31.6855
20.8277 | -35.3789 | -31.692
20.8338 | -35.3793 | -31.6986
20.8399 | -35.3799 | -31.7053
20.8461 | -35.3804 | -31.7117
20.8522 | -35.3809 | -31.7182
20.8583 | -35.3814 | -31.7247
20.8644 | -35.3819 | -31.7312
20.8705 | -35.3824 | -31.7378
20.8766 | -35.3829 | -31.7447
20.8827 | -35.3834 | -31.7512
20.8888 | -35.3839 | -31.7578
20.8949 | -35.3844 | -31.7646
20.901 | -35.3849 | -31.7712
20.9071 | -35.3853 | -31.778
20.9132 | -35.3858 | -31.7846
20.9193 | -35.3863 | -31.7912
20.9254 | -35.3869 | -31.7978
20.9315 | -35.3873 | -31.8041
20.9376 | -35.3878 | -31.8107
20.9437 | -35.3883 | -31.8175
20.9498 | -35.3888 | -31.8237
20.9559 | -35.3894 | -31.8305
20.962 | -35.3899 | -31.8372
20.9681 | -35.3903 | -31.8436
20.9742 | -35.3908 | -31.85
20.9803 | -35.3913 | -31.8564
20.9864 | -35.3918 | -31.8634
20.9926 | -35.3923 | -31.8701
20.9987 | -35.3928 | -31.8766
21.0048 | -35.3933 | -31.8833
21.0109 | -35.3938 | -31.89
21.017 | -35.3943 | -31.8965
21.0231 | -35.3948 | -31.9031
21.0292 | -35.3954 | -31.9097
21.0353 | -35.3958 | -31.9161
21.0414 | -35.3963 | -31.9227
21.0475 | -35.3969 | -31.9294
21.0536 | -35.3974 | -31.9358
21.0597 | -35.3979 | -31.9421
21.0658 | -35.3983 | -31.9488
21.0719 | -35.3988 | -31.9557
21.078 | -35.3993 | -31.9623
21.0841 | -35.3997 | -31.9688
21.0902 | -35.4002 | -31.9756
21.0963 | -35.4007 | -31.982
21.1024 | -35.4012 | -31.9885
21.1085 | -35.4017 | -31.9947
21.1146 | -35.4022 | -32.0014
21.1207 | -35.4027 | -32.008
21.1268 | -35.4032 | -32.0145
21.133 | -35.4037 | -32.0213
21.1391 | -35.4043 | -32.0275
21.1452 | -35.4048 | -32.0341
21.1513 | -35.4053 | -32.0405
21.1574 | -35.4058 | -32.0473
21.1635 | -35.4062 | -32.0537
21.1696 | -35.4067 | -32.0602
21.1757 | -35.4072 | -32.0664
21.1818 | -35.4077 | -32.0731
21.1879 | -35.4082 | -32.0796
21.194 | -35.4087 | -32.0859
21.2001 | -35.4092 | -32.0922
21.2062 | -35.4096 | -32.0988
21.2123 | -35.4102 | -32.1048
21.2184 | -35.4107 | -32.1114
21.2245 | -35.4111 | -32.1179
21.2306 | -35.4116 | -32.1246
21.2367 | -35.4121 | -32.1308
21.2428 | -35.4126 | -32.1375
21.2489 | -35.413 | -32.1439
21.255 | -35.4135 | -32.1505
21.2611 | -35.414 | -32.157
21.2672 | -35.4145 | -32.1635
21.2733 | -35.4149 | -32.17
21.2795 | -35.4154 | -32.1762
21.2856 | -35.4159 | -32.183
21.2917 | -35.4164 | -32.1897
21.2978 | -35.4168 | -32.1962
21.3039 | -35.4173 | -32.2027
21.31 | -35.4178 | -32.2091
21.3161 | -35.4182 | -32.2154
21.3222 | -35.4188 | -32.2218
21.3283 | -35.4192 | -32.2283
21.3344 | -35.4197 | -32.2345
21.3405 | -35.4202 | -32.2412
21.3466 | -35.4206 | -32.2478
21.3527 | -35.4211 | -32.2539
21.3588 | -35.4216 | -32.2602
21.3649 | -35.4221 | -32.2667
21.371 | -35.4226 | -32.2729
21.3771 | -35.4231 | -32.2792
21.3832 | -35.4235 | -32.2855
21.3893 | -35.424 | -32.2913
21.3954 | -35.4244 | -32.2977
21.4015 | -35.4249 | -32.3037
21.4076 | -35.4254 | -32.3101
21.4137 | -35.4259 | -32.3168
21.4199 | -35.4264 | -32.3231
21.426 | -35.4268 | -32.3294
21.4321 | -35.4273 | -32.3358
21.4382 | -35.4278 | -32.3421
21.4443 | -35.4282 | -32.3483
21.4504 | -35.4287 | -32.3547
21.4565 | -35.4291 | -32.3608
21.4626 | -35.4297 | -32.3672
21.4687 | -35.4301 | -32.3736
21.4748 | -35.4307 | -32.3801
21.4809 | -35.4311 | -32.3866
21.487 | -35.4316 | -32.393
21.4931 | -35.4321 | -32.3988
21.4992 | -35.4326 | -32.405
21.5053 | -35.4331 | -32.4115
21.5114 | -35.4336 | -32.4178
21.5175 | -35.4341 | -32.4242
21.5236 | -35.4345 | -32.4306
21.5297 | -35.435 | -32.4367
21.5358 | -35.4355 | -32.4431
21.5419 | -35.4359 | -32.4497
21.548 | -35.4364 | -32.4562
21.5541 | -35.4369 | -32.4628
21.5602 | -35.4374 | -32.4694
21.5664 | -35.4378 | -32.4757
21.5725 | -35.4383 | -32.4819
21.5786 | -35.4387 | -32.4884
21.5847 | -35.4391 | -32.4948
21.5908 | -35.4397 | -32.5012
21.5969 | -35.4401 | -32.5071
21.603 | -35.4407 | -32.5137
21.6091 | -35.4412 | -32.5205
21.6152 | -35.4417 | -32.5269
21.6213 | -35.4421 | -32.5335
21.6274 | -35.4426 | -32.5399
21.6335 | -35.4431 | -32.5463
21.6396 | -35.4436 | -32.5528
21.6457 | -35.4441 | -32.5594
21.6518 | -35.4445 | -32.5658
21.6579 | -35.445 | -32.5719
21.664 | -35.4455 | -32.5787
21.6701 | -35.446 | -32.5848
21.6762 | -35.4465 | -32.5913
21.6823 | -35.447 | -32.5979
21.6884 | -35.4474 | -32.6044
21.6945 | -35.4479 | -32.6107
21.7006 | -35.4484 | -32.6174
21.7068 | -35.4488 | -32.6234
21.7129 | -35.4493 | -32.6294
21.719 | -35.4498 | -32.6362
21.7251 | -35.4502 | -32.6426
21.7312 | -35.4507 | -32.6488
21.7373 | -35.4512 | -32.6552
21.7434 | -35.4516 | -32.6619
21.7495 | -35.4521 | -32.6685
21.7556 | -35.4526 | -32.6749
21.7617 | -35.453 | -32.6815
21.7678 | -35.4535 | -32.6881
21.7739 | -35.4541 | -32.6944
21.78 | -35.4545 | -32.7012
21.7861 | -35.455 | -32.7074
21.7922 | -35.4555 | -32.7139
21.7983 | -35.4559 | -32.7204
21.8044 | -35.4565 | -32.7272
21.8105 | -35.457 | -32.7333
21.8166 | -35.4575 | -32.7396
21.8227 | -35.4579 | -32.7458
21.8288 | -35.4584 | -32.752
21.8349 | -35.4588 | -32.7585
21.841 | -35.4593 | -32.7648
21.8471 | -35.4598 | -32.7711
21.8533 | -35.4603 | -32.7771
21.8594 | -35.4607 | -32.7837
21.8655 | -35.4612 | -32.7903
21.8716 | -35.4617 | -32.7969
21.8777 | -35.4621 | -32.8031
21.8838 | -35.4627 | -32.8096
21.8899 | -35.4631 | -32.8163
21.896 | -35.4636 | -32.8228
21.9021 | -35.464 | -32.8293
21.9082 | -35.4644 | -32.8356
21.9143 | -35.4649 | -32.8421
21.9204 | -35.4654 | -32.8485
21.9265 | -35.4658 | -32.855
21.9326 | -35.4663 | -32.8617
21.9387 | -35.4668 | -32.8683
21.9448 | -35.4673 | -32.8745
21.9509 | -35.4678 | -32.8805
21.957 | -35.4683 | -32.8867
21.9631 | -35.4687 | -32.8935
21.9692 | -35.4692 | -32.9001
21.9753 | -35.4696 | -32.9065
21.9814 | -35.4701 | -32.9128
21.9875 | -35.4705 | -32.9193
21.9937 | -35.471 | -32.9259
21.9998 | -35.4715 | -32.9322
22.0059 | -35.4719 | -32.9388
22.012 | -35.4724 | -32.9451
22.0181 | -35.4729 | -32.9513
22.0242 | -35.4734 | -32.9576
22.0303 | -35.4739 | -32.9639
22.0364 | -35.4743 | -32.9701
22.0425 | -35.4748 | -32.9763
22.0486 | -35.4753 | -32.9824
22.0547 | -35.4757 | -32.9886
22.0608 | -35.4762 | -32.9949
22.0669 | -35.4766 | -33.0011
22.073 | -35.4771 | -33.0071
22.0791 | -35.4776 | -33.0138
22.0852 | -35.478 | -33.02
22.0913 | -35.4784 | -33.0261
22.0974 | -35.4789 | -33.0321
22.1035 | -35.4793 | -33.0383
22.1096 | -35.4798 | -33.0447
22.1157 | -35.4803 | -33.051
22.1218 | -35.4807 | -33.0575
22.1279 | -35.4812 | -33.0633
22.134 | -35.4818 | -33.0695
22.1402 | -35.4822 | -33.0757
22.1463 | -35.4826 | -33.0819
22.1524 | -35.4831 | -33.0885
22.1585 | -35.4835 | -33.0951
22.1646 | -35.484 | -33.101
22.1707 | -35.4845 | -33.1077
22.1768 | -35.485 | -33.1141
22.1829 | -35.4854 | -33.1204
22.189 | -35.4859 | -33.127
22.1951 | -35.4863 | -33.1328
22.2012 | -35.4868 | -33.1394
22.2073 | -35.4873 | -33.1456
22.2134 | -35.4878 | -33.1517
22.2195 | -35.4883 | -33.1578
22.2256 | -35.4887 | -33.1641
22.2317 | -35.4891 | -33.1703
22.2378 | -35.4896 | -33.1764
22.2439 | -35.4901 | -33.1826
22.25 | -35.4905 | -33.189
22.2561 | -35.4911 | -33.1951
22.2622 | -35.4915 | -33.2016
22.2683 | -35.4919 | -33.2078
22.2744 | -35.4924 | -33.2139
22.2806 | -35.4928 | -33.2203
22.2867 | -35.4933 | -33.2263
22.2928 | -35.4938 | -33.2326
22.2989 | -35.4942 | -33.2388
22.305 | -35.4946 | -33.2453
22.3111 | -35.4951 | -33.2519
22.3172 | -35.4956 | -33.2579
22.3233 | -35.496 | -33.2641
22.3294 | -35.4965 | -33.2702
22.3355 | -35.497 | -33.2765
22.3416 | -35.4974 | -33.2826
22.3477 | -35.4979 | -33.2889
22.3538 | -35.4983 | -33.2951
22.3599 | -35.4988 | -33.3012
22.366 | -35.4993 | -33.3074
22.3721 | -35.4997 | -33.3135
22.3782 | -35.5002 | -33.32
22.3843 | -35.5006 | -33.3263
22.3904 | -35.5012 | -33.3325
22.3965 | -35.5016 | -33.3388
22.4026 | -35.5021 | -33.3449
22.4087 | -35.5025 | -33.3513
22.4148 | -35.5029 | -33.3575
22.4209 | -35.5034 | -33.364
22.4271 | -35.5038 | -33.3705
22.4332 | -35.5043 | -33.3767
22.4393 | -35.5047 | -33.3829
22.4454 | -35.5052 | -33.3886
22.4515 | -35.5056 | -33.3948
22.4576 | -35.506 | -33.401
22.4637 | -35.5065 | -33.4073
22.4698 | -35.507 | -33.4135
22.4759 | -35.5074 | -33.4197
22.482 | -35.5079 | -33.4259
22.4881 | -35.5084 | -33.4323
22.4942 | -35.5088 | -33.4384
22.5003 | -35.5093 | -33.4447
22.5064 | -35.5097 | -33.451
22.5125 | -35.5101 | -33.4571
22.5186 | -35.5106 | -33.4632
22.5247 | -35.511 | -33.4695
22.5308 | -35.5115 | -33.4758
22.5369 | -35.5119 | -33.4821
22.543 | -35.5124 | -33.4882
22.5491 | -35.5128 | -33.4941
22.5552 | -35.5132 | -33.5006
22.5613 | -35.5137 | -33.5071
22.5675 | -35.5141 | -33.5135
22.5736 | -35.5145 | -33.5199
22.5797 | -35.515 | -33.5261
22.5858 | -35.5155 | -33.5321
22.5919 | -35.5159 | -33.5383
22.598 | -35.5163 | -33.5446
22.6041 | -35.5168 | -33.5507
22.6102 | -35.5172 | -33.5572
22.6163 | -35.5177 | -33.5634
22.6224 | -35.5181 | -33.5697
22.6285 | -35.5186 | -33.5758
22.6346 | -35.519 | -33.5818
22.6407 | -35.5194 | -33.5881
22.6468 | -35.5199 | -33.5943
22.6529 | -35.5204 | -33.6003
22.659 | -35.5208 | -33.6063
22.6651 | -35.5212 | -33.6125
22.6712 | -35.5217 | -33.6186
22.6773 | -35.5221 | -33.6245
22.6834 | -35.5226 | -33.6307
22.6895 | -35.523 | -33.637
22.6956 | -35.5235 | -33.6433
22.7017 | -35.5239 | -33.6497
22.7079 | -35.5244 | -33.6561
22.714 | -35.5249 | -33.6624
22.7201 | -35.5253 | -33.6684
22.7262 | -35.5257 | -33.6747
22.7323 | -35.5262 | -33.6811
22.7384 | -35.5266 | -33.6874
22.7445 | -35.527 | -33.6938
22.7506 | -35.5275 | -33.6999
22.7567 | -35.528 | -33.706
22.7628 | -35.5284 | -33.7121
22.7689 | -35.5289 | -33.7183
22.775 | -35.5293 | -33.7246
22.7811 | -35.5298 | -33.7309
22.7872 | -35.5302 | -33.737
22.7933 | -35.5307 | -33.7428
22.7994 | -35.5311 | -33.7491
22.8055 | -35.5316 | -33.7554
22.8116 | -35.5321 | -33.7617
22.8177 | -35.5325 | -33.768
22.8238 | -35.5329 | -33.7745
22.8299 | -35.5334 | -33.7806
22.836 | -35.5339 | -33.7869
22.8421 | -35.5343 | -33.7931
22.8482 | -35.5347 | -33.7993
22.8544 | -35.5352 | -33.8057
22.8605 | -35.5356 | -33.8119
22.8666 | -35.5361 | -33.8182
22.8727 | -35.5365 | -33.8244
22.8788 | -35.5369 | -33.8306
22.8849 | -35.5374 | -33.8366
22.891 | -35.5378 | -33.8428
22.8971 | -35.5383 | -33.849
22.9032 | -35.5387 | -33.8551
22.9093 | -35.5392 | -33.861
22.9154 | -35.5397 | -33.8673
22.9215 | -35.5401 | -33.8734
22.9276 | -35.5405 | -33.8796
22.9337 | -35.541 | -33.8854
22.9398 | -35.5415 | -33.8917
22.9459 | -35.5419 | -33.8979
22.952 | -35.5424 | -33.9041
22.9581 | -35.5428 | -33.9104
22.9642 | -35.5432 | -33.9166
22.9703 | -35.5437 | -33.9225
22.9764 | -35.5442 | -33.9289
22.9825 | -35.5446 | -33.9351
22.9886 | -35.5451 | -33.9411
22.9948 | -35.5455 | -33.9476
23.0009 | -35.5459 | -33.9535
23.007 | -35.5464 | -33.9598
23.0131 | -35.5468 | -33.9658
23.0192 | -35.5473 | -33.9718
23.0253 | -35.5477 | -33.9781
23.0314 | -35.5482 | -33.984
23.0375 | -35.5487 | -33.9903
23.0436 | -35.5491 | -33.9966
23.0497 | -35.5495 | -34.0028
23.0558 | -35.55 | -34.0089
23.0619 | -35.5505 | -34.0152
23.068 | -35.5509 | -34.0215
23.0741 | -35.5514 | -34.0275
23.0802 | -35.5518 | -34.0336
23.0863 | -35.5522 | -34.0397
23.0924 | -35.5526 | -34.0459
23.0985 | -35.5531 | -34.052
23.1046 | -35.5535 | -34.058
23.1107 | -35.554 | -34.0643
23.1168 | -35.5544 | -34.0703
23.1229 | -35.5548 | -34.0763
23.129 | -35.5553 | -34.0827
23.1351 | -35.5557 | -34.089
23.1413 | -35.5562 | -34.0952
23.1474 | -35.5567 | -34.1011
23.1535 | -35.5571 | -34.1072
23.1596 | -35.5575 | -34.1129
23.1657 | -35.558 | -34.1189
23.1718 | -35.5584 | -34.1251
23.1779 | -35.5588 | -34.1315
23.184 | -35.5593 | -34.1372
23.1901 | -35.5597 | -34.1434
23.1962 | -35.5602 | -34.1497
23.2023 | -35.5606 | -34.1556
23.2084 | -35.5611 | -34.1618
23.2145 | -35.5615 | -34.1678
23.2206 | -35.5619 | -34.174
23.2267 | -35.5624 | -34.1802
23.2328 | -35.5628 | -34.186
23.2389 | -35.5633 | -34.1924
23.245 | -35.5637 | -34.1986
23.2511 | -35.5642 | -34.2048
23.2572 | -35.5646 | -34.2108
23.2633 | -35.5651 | -34.2169
23.2694 | -35.5655 | -34.2231
23.2755 | -35.5659 | -34.2294
23.2817 | -35.5664 | -34.2354
23.2878 | -35.5668 | -34.2414
23.2939 | -35.5673 | -34.2475
23.3 | -35.5677 | -34.2538
23.3061 | -35.5682 | -34.26
23.3122 | -35.5686 | -34.2661
23.3183 | -35.5691 | -34.2723
23.3244 | -35.5695 | -34.2784
23.3305 | -35.5699 | -34.2845
23.3366 | -35.5703 | -34.2908
23.3427 | -35.5708 | -34.2966
23.3488 | -35.5713 | -34.3028
23.3549 | -35.5717 | -34.309
23.361 | -35.5722 | -34.3151
23.3671 | -35.5727 | -34.3214
23.3732 | -35.5731 | -34.3275
23.3793 | -35.5736 | -34.3334
23.3854 | -35.574 | -34.3397
23.3915 | -35.5745 | -34.3458
23.3976 | -35.575 | -34.352
23.4037 | -35.5754 | -34.3578
23.4098 | -35.5758 | -34.3638
23.4159 | -35.5763 | -34.3698
23.422 | -35.5768 | -34.3759
23.4282 | -35.5772 | -34.3819
23.4343 | -35.5777 | -34.388
23.4404 | -35.5781 | -34.3943
23.4465 | -35.5785 | -34.4001
23.4526 | -35.5789 | -34.4065
23.4587 | -35.5794 | -34.4124
23.4648 | -35.5798 | -34.4183
23.4709 | -35.5802 | -34.4244
23.477 | -35.5805 | -34.4305
23.4831 | -35.581 | -34.4364
23.4892 | -35.5815 | -34.4427
23.4953 | -35.5819 | -34.4489
23.5014 | -35.5823 | -34.4549
23.5075 | -35.5827 | -34.4609
23.5136 | -35.5832 | -34.467
23.5197 | -35.5836 | -34.4732
23.5258 | -35.584 | -34.4792
23.5319 | -35.5844 | -34.4851
23.538 | -35.5849 | -34.4911
23.5441 | -35.5852 | -34.4972
23.5502 | -35.5857 | -34.5035
23.5563 | -35.5861 | -34.5092
23.5624 | -35.5865 | -34.5155
23.5686 | -35.5869 | -34.5214
23.5747 | -35.5874 | -34.5273
23.5808 | -35.5878 | -34.5336
23.5869 | -35.5882 | -34.5396
23.593 | -35.5887 | -34.5456
23.5991 | -35.5892 | -34.5517
23.6052 | -35.5896 | -34.5577
23.6113 | -35.5901 | -34.564
23.6174 | -35.5905 | -34.5699
23.6235 | -35.591 | -34.5758
23.6296 | -35.5914 | -34.5821
23.6357 | -35.5918 | -34.5884
23.6418 | -35.5922 | -34.5943
23.6479 | -35.5927 | -34.6007
23.654 | -35.5931 | -34.6069
23.6601 | -35.5935 | -34.6128
23.6662 | -35.594 | -34.619
23.6723 | -35.5944 | -34.6251
23.6784 | -35.5949 | -34.631
23.6845 | -35.5953 | -34.6371
23.6906 | -35.5958 | -34.6433
23.6967 | -35.5962 | -34.6496
23.7028 | -35.5966 | -34.6557
23.7089 | -35.597 | -34.6618
23.7151 | -35.5975 | -34.668
23.7212 | -35.598 | -34.6741
23.7273 | -35.5984 | -34.6803
23.7334 | -35.5988 | -34.6864
23.7395 | -35.5993 | -34.6925
23.7456 | -35.5997 | -34.6984
23.7517 | -35.6002 | -34.7045
23.7578 | -35.6006 | -34.7107
23.7639 | -35.6011 | -34.7168
23.77 | -35.6015 | -34.7228
23.7761 | -35.6019 | -34.7291
23.7822 | -35.6024 | -34.7351
23.7883 | -35.6028 | -34.7413
23.7944 | -35.6032 | -34.7475
23.8005 | -35.6037 | -34.7537
23.8066 | -35.6042 | -34.7597
23.8127 | -35.6046 | -34.7659
23.8188 | -35.605 | -34.7721
23.8249 | -35.6054 | -34.7782
23.831 | -35.6059 | -34.7845
23.8371 | -35.6063 | -34.7905
23.8432 | -35.6068 | -34.7966
23.8493 | -35.6072 | -34.8021
23.8555 | -35.6077 | -34.8084
23.8616 | -35.6081 | -34.8147
23.8677 | -35.6086 | -34.8204
23.8738 | -35.609 | -34.8266
23.8799 | -35.6094 | -34.8324
23.886 | -35.6098 | -34.8385
23.8921 | -35.6102 | -34.8449
23.8982 | -35.6107 | -34.851
23.9043 | -35.6112 | -34.857
23.9104 | -35.6116 | -34.863
23.9165 | -35.612 | -34.8688
23.9226 | -35.6124 | -34.8745
23.9287 | -35.6129 | -34.8806
23.9348 | -35.6133 | -34.8864
23.9409 | -35.6138 | -34.8926
23.947 | -35.6142 | -34.8986
23.9531 | -35.6147 | -34.9046
23.9592 | -35.6151 | -34.9107
23.9653 | -35.6155 | -34.9169
23.9714 | -35.616 | -34.9228
23.9775 | -35.6164 | -34.929
23.9836 | -35.6168 | -34.9349
23.9897 | -35.6173 | -34.9409
23.9958 | -35.6177 | -34.9474
24.002 | -35.6182 | -34.9534
24.0081 | -35.6186 | -34.9593
24.0142 | -35.619 | -34.9653
24.0203 | -35.6194 | -34.971
24.0264 | -35.6199 | -34.9771
24.0325 | -35.6203 | -34.9831
24.0386 | -35.6207 | -34.9891
24.0447 | -35.6211 | -34.9951
24.0508 | -35.6215 | -35.001
24.0569 | -35.622 | -35.0072
24.063 | -35.6225 | -35.0132
24.0691 | -35.6229 | -35.0189
24.0752 | -35.6234 | -35.0252
24.0813 | -35.6238 | -35.0311
24.0874 | -35.6242 | -35.0372
24.0935 | -35.6247 | -35.0433
24.0996 | -35.6251 | -35.0495
24.1057 | -35.6255 | -35.0554
24.1118 | -35.6259 | -35.0614
24.1179 | -35.6262 | -35.0678
24.124 | -35.6267 | -35.0736
24.1301 | -35.6271 | -35.0797
24.1362 | -35.6276 | -35.0855
24.1424 | -35.628 | -35.0915
24.1485 | -35.6285 | -35.0975
24.1546 | -35.6289 | -35.1036
24.1607 | -35.6293 | -35.1097
24.1668 | -35.6297 | -35.1157
24.1729 | -35.6302 | -35.1219
24.179 | -35.6306 | -35.1278
24.1851 | -35.631 | -35.134
24.1912 | -35.6315 | -35.1401
24.1973 | -35.6319 | -35.1461
24.2034 | -35.6323 | -35.1517
24.2095 | -35.6328 | -35.1579
24.2156 | -35.6332 | -35.164
24.2217 | -35.6336 | -35.17
24.2278 | -35.6341 | -35.1761
24.2339 | -35.6345 | -35.1819
24.24 | -35.6349 | -35.1876
24.2461 | -35.6354 | -35.1936
24.2522 | -35.6358 | -35.1995
24.2583 | -35.6363 | -35.2057
24.2644 | -35.6367 | -35.2118
24.2705 | -35.6372 | -35.2179
24.2766 | -35.6376 | -35.2237
24.2827 | -35.6381 | -35.2294
24.2889 | -35.6385 | -35.2353
24.295 | -35.639 | -35.2414
24.3011 | -35.6394 | -35.2475
24.3072 | -35.6398 | -35.2534
24.3133 | -35.6402 | -35.2591
24.3194 | -35.6407 | -35.2652
24.3255 | -35.6411 | -35.2715
24.3316 | -35.6415 | -35.2775
24.3377 | -35.6419 | -35.2836
24.3438 | -35.6423 | -35.2895
24.3499 | -35.6427 | -35.2956
24.356 | -35.6431 | -35.3015
24.3621 | -35.6435 | -35.3075
24.3682 | -35.6439 | -35.3134
24.3743 | -35.6444 | -35.3197
24.3804 | -35.6447 | -35.3253
24.3865 | -35.6452 | -35.3314
24.3926 | -35.6456 | -35.3375
24.3987 | -35.646 | -35.3434
24.4048 | -35.6464 | -35.3491
24.4109 | -35.6468 | -35.3548
24.417 | -35.6472 | -35.3607
24.4231 | -35.6477 | -35.367
24.4293 | -35.6481 | -35.3728
24.4354 | -35.6486 | -35.3785
24.4415 | -35.649 | -35.3844
24.4476 | -35.6495 | -35.3901
24.4537 | -35.6499 | -35.3958
24.4598 | -35.6504 | -35.402
24.4659 | -35.6508 | -35.4082
24.472 | -35.6512 | -35.4142
24.4781 | -35.6517 | -35.4202
24.4842 | -35.6521 | -35.4262
24.4903 | -35.6525 | -35.432
24.4964 | -35.6529 | -35.4378
24.5025 | -35.6533 | -35.4438
24.5086 | -35.6537 | -35.4497
24.5147 | -35.6542 | -35.4554
24.5208 | -35.6547 | -35.4609
24.5269 | -35.6551 | -35.4665
24.533 | -35.6555 | -35.4724
24.5391 | -35.656 | -35.4786
24.5452 | -35.6564 | -35.4845
24.5513 | -35.6569 | -35.4901
24.5574 | -35.6573 | -35.4962
24.5635 | -35.6577 | -35.502
24.5696 | -35.6582 | -35.508
24.5758 | -35.6586 | -35.5139
24.5819 | -35.6591 | -35.5199
24.588 | -35.6595 | -35.5259
24.5941 | -35.6599 | -35.5317
24.6002 | -35.6603 | -35.5377
24.6063 | -35.6608 | -35.5436
24.6124 | -35.6612 | -35.5495
24.6185 | -35.6616 | -35.5556
24.6246 | -35.6621 | -35.5611
24.6307 | -35.6625 | -35.5671
24.6368 | -35.663 | -35.5728
24.6429 | -35.6634 | -35.5787
24.649 | -35.6638 | -35.5844
24.6551 | -35.6643 | -35.5902
24.6612 | -35.6647 | -35.596
24.6673 | -35.6652 | -35.6021
24.6734 | -35.6656 | -35.608
24.6795 | -35.666 | -35.614
24.6856 | -35.6665 | -35.6201
24.6917 | -35.6669 | -35.6258
24.6978 | -35.6673 | -35.6313
24.7039 | -35.6677 | -35.6373
24.71 | -35.6682 | -35.6435
24.7162 | -35.6686 | -35.6494
24.7223 | -35.669 | -35.6554
24.7284 | -35.6694 | -35.6609
24.7345 | -35.6699 | -35.667
24.7406 | -35.6703 | -35.673
24.7467 | -35.6707 | -35.6789
24.7528 | -35.6712 | -35.6846
24.7589 | -35.6716 | -35.6903
24.765 | -35.672 | -35.696
24.7711 | -35.6725 | -35.7019
24.7772 | -35.673 | -35.7077
24.7833 | -35.6734 | -35.7136
24.7894 | -35.6738 | -35.7198
24.7955 | -35.6743 | -35.7258
24.8016 | -35.6747 | -35.7316
24.8077 | -35.6751 | -35.7375
24.8138 | -35.6755 | -35.7431
24.8199 | -35.6759 | -35.7491
24.826 | -35.6764 | -35.7549
24.8321 | -35.6768 | -35.7607
24.8382 | -35.6772 | -35.7664
24.8443 | -35.6776 | -35.7721
24.8504 | -35.678 | -35.7781
24.8565 | -35.6784 | -35.7839
24.8627 | -35.6789 | -35.7899
24.8688 | -35.6793 | -35.7955
24.8749 | -35.6797 | -35.8014
24.881 | -35.6801 | -35.8073
24.8871 | -35.6806 | -35.813
24.8932 | -35.6811 | -35.8187
24.8993 | -35.6816 | -35.8244
24.9054 | -35.682 | -35.8303
24.9115 | -35.6824 | -35.8363
24.9176 | -35.6828 | -35.8421
24.9237 | -35.6833 | -35.8481
24.9298 | -35.6837 | -35.8539
24.9359 | -35.6841 | -35.8597
24.942 | -35.6846 | -35.8652
24.9481 | -35.685 | -35.8712
24.9542 | -35.6854 | -35.877
24.9603 | -35.6859 | -35.883
24.9664 | -35.6863 | -35.8887
24.9725 | -35.6867 | -35.8948
24.9786 | -35.6871 | -35.9003
24.9847 | -35.6875 | -35.9061
24.9908 | -35.688 | -35.9118
24.9969 | -35.6884 | -35.9177
25.0031 | -35.6888 | -35.9236
25.0092 | -35.6892 | -35.9294
25.0153 | -35.6897 | -35.935
25.0214 | -35.6901 | -35.941
25.0275 | -35.6905 | -35.9469
25.0336 | -35.691 | -35.9527
25.0397 | -35.6914 | -35.9588
25.0458 | -35.6918 | -35.9647
25.0519 | -35.6922 | -35.9708
25.058 | -35.6926 | -35.9765
25.0641 | -35.6931 | -35.9823
25.0702 | -35.6935 | -35.9881
25.0763 | -35.6939 | -35.9938
25.0824 | -35.6944 | -35.9994
25.0885 | -35.6948 | -36.0052
25.0946 | -35.6952 | -36.011
25.1007 | -35.6956 | -36.0167
25.1068 | -35.6961 | -36.0223
25.1129 | -35.6965 | -36.0283
25.119 | -35.6969 | -36.0336
25.1251 | -35.6973 | -36.0393
25.1312 | -35.6977 | -36.0452
25.1373 | -35.6981 | -36.0511
25.1435 | -35.6986 | -36.0569
25.1496 | -35.699 | -36.0626
25.1557 | -35.6994 | -36.0685
25.1618 | -35.6999 | -36.0745
25.1679 | -35.7003 | -36.0803
25.174 | -35.7007 | -36.0862
25.1801 | -35.7011 | -36.0921
25.1862 | -35.7015 | -36.0977
25.1923 | -35.7019 | -36.1036
25.1984 | -35.7024 | -36.1093
25.2045 | -35.7028 | -36.1152
25.2106 | -35.7032 | -36.1211
25.2167 | -35.7036 | -36.1269
25.2228 | -35.704 | -36.1324
25.2289 | -35.7044 | -36.1382
25.235 | -35.7049 | -36.144
25.2411 | -35.7053 | -36.1499
25.2472 | -35.7057 | -36.1556
25.2533 | -35.7061 | -36.1614
25.2594 | -35.7064 | -36.1667
25.2655 | -35.7068 | -36.1725
25.2716 | -35.7073 | -36.1782
25.2777 | -35.7077 | -36.1841
25.2838 | -35.7081 | -36.19
25.29 | -35.7086 | -36.1955
25.2961 | -35.709 | -36.2012
25.3022 | -35.7094 | -36.207
25.3083 | -35.7098 | -36.2127
25.3144 | -35.7102 | -36.2184
25.3205 | -35.7106 | -36.2243
25.3266 | -35.711 | -36.23
25.3327 | -35.7114 | -36.2357
25.3388 | -35.7118 | -36.2415
25.3449 | -35.7122 | -36.2473
25.351 | -35.7126 | -36.2529
25.3571 | -35.713 | -36.2587
25.3632 | -35.7134 | -36.2642
25.3693 | -35.7139 | -36.27
25.3754 | -35.7143 | -36.2756
25.3815 | -35.7147 | -36.2813
25.3876 | -35.7151 | -36.2872
25.3937 | -35.7156 | -36.2929
25.3998 | -35.716 | -36.2985
25.4059 | -35.7164 | -36.3045
25.412 | -35.7168 | -36.3103
25.4181 | -35.7173 | -36.316
25.4242 | -35.7176 | -36.3218
25.4304 | -35.7181 | -36.3275
25.4365 | -35.7185 | -36.333
25.4426 | -35.7189 | -36.3389
25.4487 | -35.7193 | -36.3449
25.4548 | -35.7197 | -36.3506
25.4609 | -35.7202 | -36.3565
25.467 | -35.7206 | -36.3621
25.4731 | -35.7211 | -36.3677
25.4792 | -35.7215 | -36.3734
25.4853 | -35.7219 | -36.3792
25.4914 | -35.7223 | -36.3847
25.4975 | -35.7227 | -36.3903
25.5036 | -35.7232 | -36.396
25.5097 | -35.7236 | -36.4017
25.5158 | -35.724 | -36.4073
25.5219 | -35.7244 | -36.4129
25.528 | -35.7248 | -36.4187
25.5341 | -35.7253 | -36.4244
25.5402 | -35.7257 | -36.4303
25.5463 | -35.7261 | -36.4361
25.5524 | -35.7265 | -36.4419
25.5585 | -35.7269 | -36.4476
25.5646 | -35.7274 | -36.4535
25.5707 | -35.7277 | -36.4593
25.5769 | -35.7282 | -36.4648
25.583 | -35.7286 | -36.4708
25.5891 | -35.7291 | -36.4765
25.5952 | -35.7295 | -36.4822
25.6013 | -35.7299 | -36.4878
25.6074 | -35.7303 | -36.4934
25.6135 | -35.7307 | -36.499
25.6196 | -35.7311 | -36.5049
25.6257 | -35.7315 | -36.5104
25.6318 | -35.7319 | -36.5159
25.6379 | -35.7323 | -36.5215
25.644 | -35.7328 | -36.527
25.6501 | -35.7332 | -36.5326
25.6562 | -35.7336 | -36.5381
25.6623 | -35.7341 | -36.5436
25.6684 | -35.7345 | -36.5494
25.6745 | -35.7349 | -36.5551
25.6806 | -35.7354 | -36.5605
25.6867 | -35.7358 | -36.5663
25.6928 | -35.7362 | -36.5719
25.6989 | -35.7366 | -36.5777
25.705 | -35.7369 | -36.5835
25.7111 | -35.7373 | -36.5892
25.7173 | -35.7377 | -36.5945
25.7234 | -35.7381 | -36.6002
25.7295 | -35.7385 | -36.6059
25.7356 | -35.7389 | -36.6118
25.7417 | -35.7393 | -36.6172
25.7478 | -35.7397 | -36.6227
25.7539 | -35.7401 | -36.6283
25.76 | -35.7405 | -36.6339
25.7661 | -35.7409 | -36.6396
25.7722 | -35.7413 | -36.6451
25.7783 | -35.7417 | -36.6506
25.7844 | -35.7421 | -36.6561
25.7905 | -35.7425 | -36.6619
25.7966 | -35.7429 | -36.6675
25.8027 | -35.7433 | -36.6731
25.8088 | -35.7437 | -36.6787
25.8149 | -35.7441 | -36.6842
25.821 | -35.7445 | -36.6897
25.8271 | -35.7449 | -36.6955
25.8332 | -35.7454 | -36.7011
25.8393 | -35.7458 | -36.7067
25.8454 | -35.7462 | -36.7121
25.8515 | -35.7466 | -36.7178
25.8576 | -35.747 | -36.7232
25.8638 | -35.7474 | -36.7289
25.8699 | -35.7478 | -36.7347
25.876 | -35.7482 | -36.7405
25.8821 | -35.7486 | -36.7462
25.8882 | -35.7491 | -36.7514
25.8943 | -35.7495 | -36.7572
25.9004 | -35.7499 | -36.7628
25.9065 | -35.7503 | -36.7685
25.9126 | -35.7507 | -36.774
25.9187 | -35.7511 | -36.7795
25.9248 | -35.7515 | -36.7851
25.9309 | -35.7519 | -36.7909
25.937 | -35.7524 | -36.7964
25.9431 | -35.7528 | -36.8022
25.9492 | -35.7532 | -36.8077
25.9553 | -35.7536 | -36.8135
25.9614 | -35.754 | -36.8189
25.9675 | -35.7544 | -36.8243
25.9736 | -35.7548 | -36.83
25.9797 | -35.7552 | -36.8356
25.9858 | -35.7556 | -36.841
25.9919 | -35.756 | -36.8469
25.998 | -35.7564 | -36.8525
26.0042 | -35.7568 | -36.8578
26.0103 | -35.7572 | -36.8635
26.0164 | -35.7576 | -36.869
26.0225 | -35.758 | -36.8744
26.0286 | -35.7584 | -36.8802
26.0347 | -35.7588 | -36.8854
26.0408 | -35.7592 | -36.891
26.0469 | -35.7596 | -36.8964
26.053 | -35.7601 | -36.9022
26.0591 | -35.7605 | -36.9079
26.0652 | -35.7609 | -36.9132
26.0713 | -35.7613 | -36.9187
26.0774 | -35.7617 | -36.9243
26.0835 | -35.7621 | -36.9299
26.0896 | -35.7625 | -36.9354
26.0957 | -35.7629 | -36.9409
26.1018 | -35.7633 | -36.946
26.1079 | -35.7637 | -36.9514
26.114 | -35.7641 | -36.9571
26.1201 | -35.7644 | -36.9629
26.1262 | -35.7648 | -36.9688
26.1323 | -35.7652 | -36.9743
26.1384 | -35.7656 | -36.98
26.1445 | -35.766 | -36.9857
26.1507 | -35.7665 | -36.9914
26.1568 | -35.7669 | -36.9969
26.1629 | -35.7672 | -37.0023
26.169 | -35.7677 | -37.0076
26.1751 | -35.768 | -37.013
26.1812 | -35.7684 | -37.0185
26.1873 | -35.7689 | -37.0241
26.1934 | -35.7693 | -37.0298
26.1995 | -35.7697 | -37.0355
26.2056 | -35.7701 | -37.0408
26.2117 | -35.7705 | -37.0463
26.2178 | -35.7709 | -37.0517
26.2239 | -35.7714 | -37.0572
26.23 | -35.7717 | -37.0629
26.2361 | -35.7721 | -37.0687
26.2422 | -35.7725 | -37.0742
26.2483 | -35.7729 | -37.0796
26.2544 | -35.7733 | -37.0851
26.2605 | -35.7737 | -37.0908
26.2666 | -35.774 | -37.0964
26.2727 | -35.7744 | -37.102
26.2788 | -35.7748 | -37.1071
26.2849 | -35.7751 | -37.1124
26.2911 | -35.7755 | -37.1179
26.2972 | -35.7759 | -37.1235
26.3033 | -35.7763 | -37.1291
26.3094 | -35.7767 | -37.1345
26.3155 | -35.7771 | -37.1399
26.3216 | -35.7775 | -37.1454
26.3277 | -35.7779 | -37.1506
26.3338 | -35.7783 | -37.1558
26.3399 | -35.7786 | -37.1613
26.346 | -35.779 | -37.1667
26.3521 | -35.7794 | -37.1723
26.3582 | -35.7798 | -37.1776
26.3643 | -35.7802 | -37.1833
26.3704 | -35.7806 | -37.1889
26.3765 | -35.7811 | -37.1946
26.3826 | -35.7815 | -37.1999
26.3887 | -35.7819 | -37.2053
26.3948 | -35.7823 | -37.211
26.4009 | -35.7826 | -37.2166
26.407 | -35.783 | -37.2221
26.4131 | -35.7834 | -37.2274
26.4192 | -35.7839 | -37.2328
26.4253 | -35.7843 | -37.2384
26.4314 | -35.7846 | -37.2438
26.4376 | -35.785 | -37.2494
26.4437 | -35.7854 | -37.2546
26.4498 | -35.7858 | -37.2602
26.4559 | -35.7862 | -37.2655
26.462 | -35.7866 | -37.2711
26.4681 | -35.7869 | -37.2769
26.4742 | -35.7873 | -37.2826
26.4803 | -35.7878 | -37.2878
26.4864 | -35.7882 | -37.2933
26.4925 | -35.7886 | -37.2989
26.4986 | -35.789 | -37.3046
26.5047 | -35.7894 | -37.3102
26.5108 | -35.7898 | -37.3157
26.5169 | -35.7902 | -37.3212
26.523 | -35.7905 | -37.3267
26.5291 | -35.7909 | -37.3319
26.5352 | -35.7913 | -37.3377
26.5413 | -35.7917 | -37.3435
26.5474 | -35.7921 | -37.3491
26.5535 | -35.7925 | -37.3549
26.5596 | -35.7928 | -37.3605
26.5657 | -35.7932 | -37.3658
26.5718 | -35.7936 | -37.3715
26.578 | -35.7939 | -37.3771
26.5841 | -35.7944 | -37.3827
26.5902 | -35.7947 | -37.3879
26.5963 | -35.7951 | -37.3933
26.6024 | -35.7955 | -37.3989
26.6085 | -35.7959 | -37.4045
26.6146 | -35.7963 | -37.4097
26.6207 | -35.7966 | -37.4154
26.6268 | -35.797 | -37.4212
26.6329 | -35.7974 | -37.4264
26.639 | -35.7978 | -37.4319
26.6451 | -35.7982 | -37.4372
26.6512 | -35.7986 | -37.4425
26.6573 | -35.7989 | -37.4482
26.6634 | -35.7993 | -37.4533
26.6695 | -35.7997 | -37.4591
26.6756 | -35.8001 | -37.4649
26.6817 | -35.8004 | -37.4704
26.6878 | -35.8008 | -37.4761
26.6939 | -35.8012 | -37.4813
26.7 | -35.8016 | -37.4864
26.7061 | -35.802 | -37.4917
26.7122 | -35.8023 | -37.4974
26.7183 | -35.8027 | -37.5027
26.7245 | -35.803 | -37.5081
26.7306 | -35.8034 | -37.5135
26.7367 | -35.8037 | -37.5185
26.7428 | -35.8041 | -37.524
26.7489 | -35.8045 | -37.5294
26.755 | -35.8048 | -37.5352
26.7611 | -35.8052 | -37.5409
26.7672 | -35.8056 | -37.5462
26.7733 | -35.806 | -37.5516
26.7794 | -35.8064 | -37.5571
26.7855 | -35.8068 | -37.5627
26.7916 | -35.8071 | -37.568
26.7977 | -35.8076 | -37.5735
26.8038 | -35.8079 | -37.579
26.8099 | -35.8083 | -37.5843
26.816 | -35.8087 | -37.5899
26.8221 | -35.8091 | -37.5951
26.8282 | -35.8095 | -37.6007
26.8343 | -35.8099 | -37.6063
26.8404 | -35.8103 | -37.6119
26.8465 | -35.8106 | -37.6174
26.8526 | -35.811 | -37.623
26.8587 | -35.8113 | -37.6284
26.8649 | -35.8118 | -37.6338
26.871 | -35.8121 | -37.6391
26.8771 | -35.8125 | -37.6447
26.8832 | -35.8129 | -37.6503
26.8893 | -35.8133 | -37.6558
26.8954 | -35.8137 | -37.6612
26.9015 | -35.8141 | -37.667
26.9076 | -35.8145 | -37.6724
26.9137 | -35.8149 | -37.6782
26.9198 | -35.8153 | -37.684
26.9259 | -35.8156 | -37.6894
26.932 | -35.8159 | -37.6943
26.9381 | -35.8163 | -37.7001
26.9442 | -35.8167 | -37.7048
26.9503 | -35.817 | -37.7107
26.9564 | -35.8174 | -37.7162
26.9625 | -35.8178 | -37.7218
26.9686 | -35.8182 | -37.7271
26.9747 | -35.8186 | -37.7324
26.9808 | -35.8189 | -37.7377
26.9869 | -35.8193 | -37.7428
26.993 | -35.8197 | -37.7481
26.9991 | -35.82 | -37.7532
27.0052 | -35.8204 | -37.7589
27.0114 | -35.8208 | -37.7643
27.0175 | -35.8212 | -37.7696
27.0236 | -35.8215 | -37.7748
27.0297 | -35.8218 | -37.7803
27.0358 | -35.8222 | -37.7856
27.0419 | -35.8226 | -37.7911
27.048 | -35.8229 | -37.7964
27.0541 | -35.8233 | -37.8019
27.0602 | -35.8236 | -37.8073
27.0663 | -35.824 | -37.8132
27.0724 | -35.8244 | -37.8184
27.0785 | -35.8248 | -37.8239
27.0846 | -35.8251 | -37.8292
27.0907 | -35.8255 | -37.8346
27.0968 | -35.8259 | -37.8401
27.1029 | -35.8263 | -37.8452
27.109 | -35.8266 | -37.8504
27.1151 | -35.827 | -37.8559
27.1212 | -35.8273 | -37.8606
27.1273 | -35.8278 | -37.8659
27.1334 | -35.8281 | -37.8714
27.1395 | -35.8285 | -37.8767
27.1456 | -35.8289 | -37.8823
27.1518 | -35.8292 | -37.8878
27.1579 | -35.8296 | -37.8929
27.164 | -35.83 | -37.8981
27.1701 | -35.8304 | -37.9035
27.1762 | -35.8307 | -37.9091
27.1823 | -35.8311 | -37.9145
27.1884 | -35.8315 | -37.9201
27.1945 | -35.8318 | -37.9254
27.2006 | -35.8322 | -37.9307
27.2067 | -35.8325 | -37.9361
27.2128 | -35.8329 | -37.9416
27.2189 | -35.8333 | -37.947
27.225 | -35.8337 | -37.9524
27.2311 | -35.834 | -37.9574
27.2372 | -35.8344 | -37.9626
27.2433 | -35.8348 | -37.9679
27.2494 | -35.8352 | -37.9732
27.2555 | -35.8356 | -37.9788
27.2616 | -35.8359 | -37.9839
27.2677 | -35.8363 | -37.9889
27.2738 | -35.8366 | -37.9941
27.2799 | -35.837 | -37.9996
27.286 | -35.8373 | -38.0049
27.2921 | -35.8377 | -38.0102
27.2983 | -35.838 | -38.0151
27.3044 | -35.8384 | -38.0203
27.3105 | -35.8388 | -38.0256
27.3166 | -35.8392 | -38.031
27.3227 | -35.8396 | -38.0362
27.3288 | -35.84 | -38.0416
27.3349 | -35.8404 | -38.0465
27.341 | -35.8407 | -38.0517
27.3471 | -35.841 | -38.0567
27.3532 | -35.8414 | -38.062
27.3593 | -35.8418 | -38.0676
27.3654 | -35.8421 | -38.0733
27.3715 | -35.8426 | -38.0782
27.3776 | -35.843 | -38.0838
27.3837 | -35.8434 | -38.0889
27.3898 | -35.8437 | -38.0942
27.3959 | -35.8441 | -38.0995
27.402 | -35.8445 | -38.1047
27.4081 | -35.8449 | -38.1099
27.4142 | -35.8453 | -38.1153
27.4203 | -35.8456 | -38.1205
27.4264 | -35.846 | -38.126
27.4325 | -35.8464 | -38.1311
27.4387 | -35.8467 | -38.1363
27.4448 | -35.8471 | -38.1416
27.4509 | -35.8474 | -38.1468
27.457 | -35.8479 | -38.1519
27.4631 | -35.8482 | -38.1574
27.4692 | -35.8485 | -38.1625
27.4753 | -35.8488 | -38.1674
27.4814 | -35.8491 | -38.1722
27.4875 | -35.8495 | -38.1777
27.4936 | -35.8499 | -38.183
27.4997 | -35.8503 | -38.1883
27.5058 | -35.8506 | -38.1938
27.5119 | -35.851 | -38.1986
27.518 | -35.8514 | -38.2038
27.5241 | -35.8518 | -38.2093
27.5302 | -35.8521 | -38.2146
27.5363 | -35.8525 | -38.2197
27.5424 | -35.8529 | -38.2245
27.5485 | -35.8532 | -38.2298
27.5546 | -35.8536 | -38.2349
27.5607 | -35.854 | -38.2404
27.5668 | -35.8543 | -38.2456
27.5729 | -35.8547 | -38.2506
27.5791 | -35.855 | -38.2558
27.5852 | -35.8554 | -38.2608
27.5913 | -35.8558 | -38.2661
27.5974 | -35.8562 | -38.2714
27.6035 | -35.8565 | -38.2769
27.6096 | -35.8569 | -38.2819
27.6157 | -35.8572 | -38.2869
27.6218 | -35.8576 | -38.2926
27.6279 | -35.8579 | -38.2979
27.634 | -35.8583 | -38.3031
27.6401 | -35.8587 | -38.3085
27.6462 | -35.859 | -38.3139
27.6523 | -35.8594 | -38.3191
27.6584 | -35.8598 | -38.3241
27.6645 | -35.8601 | -38.3295
27.6706 | -35.8605 | -38.3346
27.6767 | -35.8609 | -38.3398
27.6828 | -35.8611 | -38.3447
27.6889 | -35.8615 | -38.3499
27.695 | -35.8618 | -38.3552
27.7011 | -35.8622 | -38.3603
27.7072 | -35.8625 | -38.3658
27.7133 | -35.8629 | -38.3712
27.7194 | -35.8633 | -38.3759
27.7256 | -35.8636 | -38.3812
27.7317 | -35.864 | -38.3862
27.7378 | -35.8643 | -38.3918
27.7439 | -35.8647 | -38.397
27.75 | -35.8651 | -38.4023
27.7561 | -35.8654 | -38.4074
27.7622 | -35.8658 | -38.4127
27.7683 | -35.8661 | -38.4181
27.7744 | -35.8665 | -38.4235
27.7805 | -35.8668 | -38.4285
27.7866 | -35.8671 | -38.4336
27.7927 | -35.8674 | -38.439
27.7988 | -35.8678 | -38.4439
27.8049 | -35.8682 | -38.449
27.811 | -35.8685 | -38.4543
27.8171 | -35.8689 | -38.4596
27.8232 | -35.8692 | -38.4649
27.8293 | -35.8695 | -38.4702
27.8354 | -35.8699 | -38.4754
27.8415 | -35.8703 | -38.4804
27.8476 | -35.8707 | -38.4855
27.8537 | -35.871 | -38.4909
27.8598 | -35.8714 | -38.4959
27.866 | -35.8717 | -38.501
27.8721 | -35.8721 | -38.5066
27.8782 | -35.8724 | -38.5119
27.8843 | -35.8727 | -38.5174
27.8904 | -35.8731 | -38.5227
27.8965 | -35.8734 | -38.5276
27.9026 | -35.8737 | -38.5328
27.9087 | -35.8741 | -38.538
27.9148 | -35.8744 | -38.543
27.9209 | -35.8748 | -38.5483
27.927 | -35.8751 | -38.5535
27.9331 | -35.8754 | -38.5587
27.9392 | -35.8758 | -38.5637
27.9453 | -35.8762 | -38.5687
27.9514 | -35.8765 | -38.5738
27.9575 | -35.8769 | -38.5787
27.9636 | -35.8772 | -38.5839
27.9697 | -35.8775 | -38.5889
27.9758 | -35.8779 | -38.5941
27.9819 | -35.8783 | -38.5991
27.988 | -35.8786 | -38.6041
27.9941 | -35.879 | -38.6094
28.0002 | -35.8793 | -38.6143
28.0063 | -35.8795 | -38.6192
28.0125 | -35.8798 | -38.6243
28.0186 | -35.8801 | -38.6297
28.0247 | -35.8805 | -38.6348
28.0308 | -35.8809 | -38.6398
28.0369 | -35.8812 | -38.6447
28.043 | -35.8816 | -38.6504
28.0491 | -35.8819 | -38.6556
28.0552 | -35.8822 | -38.661
28.0613 | -35.8825 | -38.6665
28.0674 | -35.8829 | -38.6714
28.0735 | -35.8832 | -38.6764

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.5809 | -42.7204
21.4908 | -35.6347 | -35.6104
16.4948 | -34.9989 | -28.1715
12.6602 | -34.6273 | -21.8441
9.7171 | -34.3962 | -16.9001
7.4582 | -34.2444 | -12.8835
5.7244 | -34.124 | -9.6304
4.3936 | -34.049 | -7.0087
3.3722 | -34.0313 | -4.9369
2.5883 | -34.0184 | -3.3669
1.9866 | -34.0137 | -2.1149
1.5248 | -34.0158 | -1.107
1.1703 | -34.0198 | -0.44987
1 | -34.0196 | -0.1241
0.89822 | -34.0036 | 0.15098
0.68941 | -33.985 | 0.74019
0.52919 | -33.9809 | 1.1905
0.40612 | -33.9826 | 1.806
0.31171 | -33.9696 | 2.2584
0.23928 | -33.9722 | 2.9841
0.18358 | -33.9881 | 3.3938
0.14091 | -33.9854 | 4.2782
0.10819 | -34.0041 | 5.4391
0.083017 | -34.0072 | 6.9279
0.063705 | -34.0634 | 8.7952
0.048923 | -34.1297 | 11.286
0.037479 | -34.2291 | 14.4812
0.028801 | -34.3998 | 18.5182
0.022078 | -34.7078 | 23.9189
0.017023 | -35.1931 | 30.346
0.013018 | -35.9931 | 38.0029
0.010014 | -37.0972 | 45.9051

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003613 | -2.8589
0.007227 | -28.1401
0.01084 | -35.631
0.014453 | -40.9982
0.018067 | -44.6145
0.02168 | -47.035
0.025294 | -48.3721
0.028907 | -50.6559
0.03252 | -51.822
0.036134 | -52.527
0.039747 | -53.8992
0.04336 | -54.9627
0.046974 | -55.446
0.050587 | -56.754
0.054201 | -58.3854
0.057814 | -58.3792
0.061427 | -58.86
0.065041 | -59.9717
0.068654 | -61.1656
0.072267 | -62.1378
0.075881 | -62.1963
0.079494 | -60.8461
0.083107 | -62.0783
0.086721 | -62.934
0.090334 | -62.672
0.093948 | -63.1418
0.097561 | -64.0447
0.10117 | -64.3715
0.10479 | -64.5978
0.1084 | -65.0567
0.11201 | -65.5314
0.11563 | -65.9024
0.11924 | -66.9935
0.12286 | -67.3822
0.12647 | -66.4508
0.13008 | -67.0208
0.1337 | -66.9952
0.13731 | -67.3825
0.14092 | -68.3243
0.14453 | -68.7855
0.14815 | -68.1041
0.15176 | -68.1514
0.15538 | -68.9341
0.15899 | -69.2866
0.1626 | -69.7322
0.16622 | -70.2554
0.16983 | -70.2505
0.17344 | -69.7415
0.17705 | -70.7631
0.18067 | -71.4609
0.18428 | -71.1217
0.1879 | -71.0512
0.19151 | -71.4137
0.19512 | -71.4518
0.19873 | -70.8
0.20235 | -70.8623
0.20596 | -71.552
0.20958 | -72.3402
0.21319 | -72.5154
0.2168 | -72.5909
0.22042 | -72.3772
0.22403 | -72.8733
0.22764 | -73.1871
0.23126 | -73.0831
0.23487 | -73.3216
0.23848 | -74.1039
0.2421 | -73.4493
0.24571 | -73.8546
0.24932 | -74.7053
0.25294 | -75.4324
0.25655 | -75.1463
0.26016 | -75.1883
0.26378 | -74.1658
0.26739 | -74.6218
0.271 | -75.1314
0.27462 | -75.0919
0.27823 | -76.2144
0.28184 | -76.0104
0.28546 | -75.6363
0.28907 | -76.5322
0.29268 | -76.8414
0.2963 | -76.1395
0.29991 | -75.3183
0.30352 | -76.2036
0.30714 | -77.2629
0.31075 | -77.4799
0.31436 | -76.8826
0.31798 | -77.0877
0.32159 | -78.0855
0.3252 | -77.4471
0.32882 | -78.031
0.33243 | -77.4372
0.33604 | -76.3117
0.33966 | -77.2819
0.34327 | -78.4036
0.34688 | -78.7445
0.3505 | -78.6941
0.35411 | -78.2773
0.35772 | -77.9361
0.36134 | -77.1867
0.36495 | -77.6035
0.36856 | -77.7484
0.37218 | -79.2946
0.37579 | -79.7408
0.3794 | -78.9836
0.38302 | -78.405
0.38663 | -78.1663
0.39024 | -78.5439
0.39386 | -78.1933
0.39747 | -77.9879
0.40108 | -78.6367
0.4047 | -79.0738
0.40831 | -79.258
0.41192 | -79.2265
0.41554 | -80.1033
0.41915 | -80.0108
0.42276 | -79.3071
0.42638 | -79.6656
0.42999 | -79.8998
0.4336 | -79.3398
0.43722 | -80.2501
0.44083 | -79.57
0.44444 | -78.627
0.44806 | -79.2202
0.45167 | -80.8887
0.45528 | -81.1308
0.4589 | -80.8349
0.46251 | -80.1548
0.46613 | -80.3691
0.46974 | -80.345
0.47335 | -80.4393
0.47696 | -80.6632
0.48058 | -80.8655
0.48419 | -80.8018
0.4878 | -81.0557
0.49142 | -80.1083
0.49503 | -80.6484
0.49865 | -80.8519
0.50226 | -81.6199
0.50587 | -81.0565
0.50948 | -81.241
0.5131 | -82.4201
0.51671 | -82.1451
0.52033 | -80.9795
0.52394 | -80.4591
0.52755 | -80.4782
0.53116 | -80.9406
0.53478 | -80.9959
0.53839 | -80.6865
0.542 | -80.8042
0.54562 | -81.2175
0.54923 | -81.5963
0.55285 | -82.028
0.55646 | -82.0521
0.56007 | -81.607
0.56369 | -81.2158
0.5673 | -80.8315
0.57091 | -80.641
0.57453 | -80.9938
0.57814 | -82.0256
0.58175 | -82.8062
0.58537 | -82.6147
0.58898 | -82.0666
0.59259 | -82.0277
0.59621 | -82.199
0.59982 | -82.3433
0.60343 | -82.8855
0.60705 | -83.0747
0.61066 | -82.5655
0.61427 | -82.3734
0.61789 | -82.7761
0.6215 | -82.9971
0.62511 | -82.6387
0.62873 | -81.9271
0.63234 | -81.7772
0.63595 | -82.3156
0.63957 | -83.1723
0.64318 | -83.8673
0.64679 | -84.0209
0.65041 | -84.1905
0.65402 | -84.1537
0.65763 | -83.9209
0.66125 | -83.8385
0.66486 | -83.3953
0.66847 | -82.9397
0.67209 | -83.19
0.6757 | -83.661
0.67931 | -83.8494
0.68293 | -83.9984
0.68654 | -84.0408
0.69015 | -84.2575
0.69377 | -84.6123
0.69738 | -84.5045
0.70099 | -83.9672
0.70461 | -83.8565
0.70822 | -84.1023
0.71183 | -84.4144
0.71545 | -84.2569
0.71906 | -83.8738
0.72267 | -84.0376
0.72629 | -84.5003
0.7299 | -84.8236
0.73351 | -84.7697
0.73713 | -84.6146
0.74074 | -84.6666
0.74435 | -84.9816
0.74797 | -85.6435
0.75158 | -85.8861
0.75519 | -85.5401
0.75881 | -84.8533
0.76242 | -84.6367
0.76603 | -85.1179
0.76965 | -85.7954
0.77326 | -86.2076
0.77687 | -86.4227
0.78049 | -86.5513
0.7841 | -86.3199
0.78772 | -85.3608
0.79133 | -84.3225
0.79494 | -84.0618
0.79856 | -84.5227
0.80217 | -85.1699
0.80578 | -85.712
0.80939 | -85.5678
0.81301 | -85.2427
0.81662 | -85.2587
0.82024 | -85.5126
0.82385 | -85.4384
0.82746 | -85.3005
0.83108 | -85.5895
0.83469 | -85.9563
0.8383 | -86.0215
0.84191 | -85.9534
0.84553 | -85.9489
0.84914 | -85.6746
0.85276 | -85.3289
0.85637 | -85.3871
0.85998 | -85.6207
0.8636 | -86.1113
0.86721 | -86.5358
0.87082 | -86.3917
0.87443 | -86.0246
0.87805 | -85.5258
0.88166 | -85.3148
0.88528 | -85.6419
0.88889 | -86.2812
0.8925 | -87.0278
0.89612 | -87.4976
0.89973 | -87.5422
0.90334 | -87.2431
0.90696 | -86.9741
0.91057 | -87.0169
0.91418 | -86.7209
0.9178 | -86.5672
0.92141 | -87.2157
0.92502 | -87.927
0.92864 | -87.4335
0.93225 | -86.6035
0.93586 | -86.6368
0.93948 | -86.6114
0.94309 | -86.0082
0.9467 | -85.7333
0.95032 | -86.0795
0.95393 | -86.5197
0.95754 | -86.9395
0.96116 | -87.313
0.96477 | -87.0845
0.96838 | -87.3161
0.972 | -88.134
0.97561 | -88.2839
0.97922 | -87.7038
0.98284 | -87.2756
0.98645 | -87.1357
0.99006 | -86.6951
0.99368 | -86.5361
0.99729 | -86.8611
1.0009 | -87.4032
1.0045 | -87.9285
1.0081 | -88.1201
1.0117 | -87.7075
1.0154 | -87.2093
1.019 | -86.8991
1.0226 | -86.9507
1.0262 | -87.5605
1.0298 | -87.9616
1.0334 | -87.7823
1.037 | -87.7852
1.0407 | -88.3161
1.0443 | -88.7121
1.0479 | -88.7547
1.0515 | -88.322
1.0551 | -87.7586
1.0587 | -87.4901
1.0623 | -87.7045
1.0659 | -88.2493
1.0696 | -88.7259
1.0732 | -88.807
1.0768 | -88.5159
1.0804 | -88.0812
1.084 | -87.8024
1.0876 | -87.8214
1.0912 | -88.0424
1.0949 | -88.424
1.0985 | -88.9805
1.1021 | -89.4475
1.1057 | -89.5928
1.1093 | -89.5391
1.1129 | -89.4241
1.1165 | -89.1678
1.1201 | -88.7671
1.1238 | -88.3733
1.1274 | -88.1116
1.131 | -88.0488
1.1346 | -88.1489
1.1382 | -88.2429
1.1418 | -88.2586
1.1454 | -88.3898
1.1491 | -88.7962
1.1527 | -89.2661
1.1563 | -89.5115
1.1599 | -89.493
1.1635 | -89.3364
1.1671 | -89.2009
1.1707 | -89.1228
1.1743 | -89.1015
1.178 | -89.1546
1.1816 | -89.0894
1.1852 | -88.8006
1.1888 | -88.5225
1.1924 | -88.4401
1.196 | -88.5656
1.1996 | -88.8355
1.2033 | -89.117
1.2069 | -89.3331
1.2105 | -89.3992
1.2141 | -89.2326
1.2177 | -88.9383
1.2213 | -88.6941
1.2249 | -88.624
1.2285 | -88.7328
1.2322 | -88.9015
1.2358 | -89.0118
1.2394 | -88.9973
1.243 | -88.9854
1.2466 | -89.135
1.2502 | -89.2874
1.2538 | -89.2681
1.2575 | -89.293
1.2611 | -89.6086
1.2647 | -90.0396
1.2683 | -90.1217
1.2719 | -89.7847
1.2755 | -89.4656
1.2791 | -89.4691
1.2827 | -89.7699
1.2864 | -90.091
1.29 | -90.1038
1.2936 | -89.8514
1.2972 | -89.7751
1.3008 | -90.1067
1.3044 | -90.5376
1.308 | -90.6086
1.3117 | -90.4273
1.3153 | -90.1855
1.3189 | -89.8081
1.3225 | -89.4284
1.3261 | -89.2787
1.3297 | -89.3468
1.3333 | -89.5069
1.3369 | -89.6934
1.3406 | -89.8529
1.3442 | -89.843
1.3478 | -89.6675
1.3514 | -89.5178
1.355 | -89.4463
1.3586 | -89.4086
1.3622 | -89.4573
1.3659 | -89.5563
1.3695 | -89.5678
1.3731 | -89.5758
1.3767 | -89.7706
1.3803 | -90.0302
1.3839 | -90.0624
1.3875 | -89.9285
1.3911 | -89.8109
1.3948 | -89.7751
1.3984 | -89.8596
1.402 | -89.9854
1.4056 | -90.1082
1.4092 | -90.3306
1.4128 | -90.6891
1.4164 | -91.0248
1.4201 | -91.0836
1.4237 | -90.9215
1.4273 | -90.7388
1.4309 | -90.4843
1.4345 | -90.0605
1.4381 | -89.7651
1.4417 | -89.9577
1.4453 | -90.5754
1.449 | -91.1484
1.4526 | -91.1519
1.4562 | -90.6597
1.4598 | -90.2172
1.4634 | -90.0883
1.467 | -90.1915
1.4706 | -90.3612
1.4743 | -90.4977
1.4779 | -90.5488
1.4815 | -90.5128
1.4851 | -90.4671
1.4887 | -90.4301
1.4923 | -90.3211
1.4959 | -90.045
1.4995 | -89.7113
1.5032 | -89.5926
1.5068 | -89.7895
1.5104 | -90.2066
1.514 | -90.6615
1.5176 | -90.9755
1.5212 | -91.2282
1.5248 | -91.5826
1.5285 | -91.8556
1.5321 | -91.7609
1.5357 | -91.4535
1.5393 | -91.3048
1.5429 | -91.3314
1.5465 | -91.3062
1.5501 | -91.224
1.5537 | -91.188
1.5574 | -91.2562
1.561 | -91.5134
1.5646 | -91.6729
1.5682 | -91.6889
1.5718 | -91.5398
1.5754 | -91.3475
1.579 | -91.2575
1.5827 | -91.3602
1.5863 | -91.6139
1.5899 | -91.8708
1.5935 | -92.0113
1.5971 | -91.9579
1.6007 | -91.698
1.6043 | -91.3406
1.6079 | -91.0708
1.6116 | -91.015
1.6152 | -91.1523
1.6188 | -91.3615
1.6224 | -91.5237
1.626 | -91.6249
1.6296 | -91.7415
1.6332 | -91.8958
1.6369 | -92.04
1.6405 | -92.1496
1.6441 | -92.2421
1.6477 | -92.3437
1.6513 | -92.4287
1.6549 | -92.4741
1.6585 | -92.4694
1.6622 | -92.4039
1.6658 | -92.3055
1.6694 | -92.2478
1.673 | -92.2859
1.6766 | -92.3833
1.6802 | -92.4478
1.6838 | -92.4389
1.6874 | -92.3398
1.6911 | -92.1451
1.6947 | -91.868
1.6983 | -91.5673
1.7019 | -91.2978
1.7055 | -91.1225
1.7091 | -91.0741
1.7127 | -91.1496
1.7164 | -91.2917
1.72 | -91.403
1.7236 | -91.4078
1.7272 | -91.345
1.7308 | -91.3156
1.7344 | -91.3532
1.738 | -91.4142
1.7416 | -91.499
1.7453 | -91.6548
1.7489 | -91.8952
1.7525 | -92.1543
1.7561 | -92.3489
1.7597 | -92.4238
1.7633 | -92.3494
1.7669 | -92.1794
1.7706 | -91.9904
1.7742 | -91.8939
1.7778 | -91.9073
1.7814 | -91.9578
1.785 | -91.9837
1.7886 | -92.0209
1.7922 | -92.1553
1.7958 | -92.3843
1.7995 | -92.6096
1.8031 | -92.7409
1.8067 | -92.74
1.8103 | -92.7096
1.8139 | -92.719
1.8175 | -92.7376
1.8211 | -92.7068
1.8248 | -92.6361
1.8284 | -92.5973
1.832 | -92.6429
1.8356 | -92.7555
1.8392 | -92.8898
1.8428 | -93.0146
1.8464 | -93.0777
1.85 | -92.9737
1.8537 | -92.6781
1.8573 | -92.3743
1.8609 | -92.297
1.8645 | -92.496
1.8681 | -92.8003
1.8717 | -92.9554
1.8753 | -92.9161
1.879 | -92.843
1.8826 | -92.8564
1.8862 | -92.964
1.8898 | -93.1253
1.8934 | -93.275
1.897 | -93.3802
1.9006 | -93.4197
1.9042 | -93.3648
1.9079 | -93.2399
1.9115 | -93.0718
1.9151 | -92.9202
1.9187 | -92.8357
1.9223 | -92.8125
1.9259 | -92.8085
1.9295 | -92.7737
1.9332 | -92.7374
1.9368 | -92.7554
1.9404 | -92.8144
1.944 | -92.958
1.9476 | -93.2119
1.9512 | -93.5198
1.9548 | -93.7107
1.9584 | -93.7023
1.9621 | -93.6381
1.9657 | -93.6915
1.9693 | -93.8953
1.9729 | -94.1339
1.9765 | -94.2524
1.9801 | -94.2089
1.9837 | -94.0176
1.9874 | -93.7295
1.991 | -93.3991
1.9946 | -93.0909
1.9982 | -92.8625
2.0018 | -92.7201
2.0054 | -92.6516
2.009 | -92.6795
2.0126 | -92.8441
2.0163 | -93.1328
2.0199 | -93.4576
2.0235 | -93.7538
2.0271 | -94.0009
2.0307 | -94.1584
2.0343 | -94.182
2.0379 | -94.0348
2.0416 | -93.76
2.0452 | -93.4932
2.0488 | -93.3754
2.0524 | -93.4621
2.056 | -93.6843
2.0596 | -93.8984
2.0632 | -93.9955
2.0668 | -93.9935
2.0705 | -93.9849
2.0741 | -94.0091
2.0777 | -94.0424
2.0813 | -94.0528
2.0849 | -94.0379
2.0885 | -94.0368
2.0921 | -94.0556
2.0958 | -94.0913
2.0994 | -94.1225
2.103 | -94.1177
2.1066 | -94.0902
2.1102 | -94.058
2.1138 | -94.0227
2.1174 | -93.9603
2.121 | -93.8357
2.1247 | -93.6298
2.1283 | -93.3645
2.1319 | -93.0945
2.1355 | -92.8791
2.1391 | -92.7656
2.1427 | -92.7868
2.1463 | -92.9394
2.15 | -93.1887
2.1536 | -93.4641
2.1572 | -93.6745
2.1608 | -93.7804
2.1644 | -93.8164
2.168 | -93.8431
2.1716 | -93.9082
2.1752 | -94.0161
2.1789 | -94.1317
2.1825 | -94.2359
2.1861 | -94.3128
2.1897 | -94.3645
2.1933 | -94.4033
2.1969 | -94.4585
2.2005 | -94.5549
2.2042 | -94.684
2.2078 | -94.7826
2.2114 | -94.7983
2.215 | -94.7134
2.2186 | -94.5549
2.2222 | -94.378
2.2258 | -94.2613
2.2294 | -94.2115
2.2331 | -94.1883
2.2367 | -94.1573
2.2403 | -94.0987
2.2439 | -94.0389
2.2475 | -94.001
2.2511 | -93.9929
2.2547 | -93.9966
2.2584 | -94.0114
2.262 | -94.0601
2.2656 | -94.1449
2.2692 | -94.2458
2.2728 | -94.3427
2.2764 | -94.4017
2.28 | -94.4087
2.2837 | -94.3628
2.2873 | -94.2977
2.2909 | -94.2511
2.2945 | -94.2555
2.2981 | -94.2827
2.3017 | -94.2961
2.3053 | -94.2575
2.3089 | -94.1663
2.3126 | -94.0616
2.3162 | -94.0035
2.3198 | -94.0437
2.3234 | -94.186
2.327 | -94.3764
2.3306 | -94.5468
2.3342 | -94.658
2.3378 | -94.6875
2.3415 | -94.6168
2.3451 | -94.4781
2.3487 | -94.3043
2.3523 | -94.1462
2.3559 | -94.0302
2.3595 | -93.9268
2.3631 | -93.8078
2.3668 | -93.6816
2.3704 | -93.5797
2.374 | -93.5413
2.3776 | -93.6105
2.3812 | -93.8105
2.3848 | -94.1243
2.3884 | -94.4709
2.3921 | -94.7084
2.3957 | -94.7285
2.3993 | -94.5695
2.4029 | -94.3663
2.4065 | -94.195
2.4101 | -94.0952
2.4137 | -94.0687
2.4173 | -94.105
2.421 | -94.1762
2.4246 | -94.269
2.4282 | -94.3688
2.4318 | -94.4714
2.4354 | -94.5744
2.439 | -94.6607
2.4426 | -94.7326
2.4463 | -94.8116
2.4499 | -94.8624
2.4535 | -94.8293
2.4571 | -94.7061
2.4607 | -94.5734
2.4643 | -94.505
2.4679 | -94.5439
2.4715 | -94.6376
2.4752 | -94.6994
2.4788 | -94.701
2.4824 | -94.6591
2.486 | -94.5927
2.4896 | -94.517
2.4932 | -94.4268
2.4968 | -94.3106
2.5005 | -94.1925
2.5041 | -94.115
2.5077 | -94.1042
2.5113 | -94.1685
2.5149 | -94.2944
2.5185 | -94.4338
2.5221 | -94.5535
2.5257 | -94.6248
2.5294 | -94.6338
2.533 | -94.6251
2.5366 | -94.67
2.5402 | -94.8013
2.5438 | -95.0074
2.5474 | -95.2043
2.551 | -95.2911
2.5547 | -95.2395
2.5583 | -95.1087
2.5619 | -94.9914
2.5655 | -94.9467
2.5691 | -94.9458
2.5727 | -94.9281
2.5763 | -94.8665
2.5799 | -94.7605
2.5836 | -94.653
2.5872 | -94.5654
2.5908 | -94.5027
2.5944 | -94.4635
2.598 | -94.466
2.6016 | -94.5183
2.6052 | -94.6696
2.6089 | -94.8008
2.6125 | -94.9528
2.6161 | -95.0984
2.6197 | -95.1957
2.6233 | -95.2487
2.6269 | -95.2575
2.6305 | -95.2637
2.6341 | -95.2876
2.6378 | -95.3588
2.6414 | -95.4702
2.645 | -95.6084
2.6486 | -95.7502
2.6522 | -95.8494
2.6558 | -95.8763
2.6594 | -95.8055
2.6631 | -95.6716
2.6667 | -95.5239
2.6703 | -95.3948
2.6739 | -95.3053
2.6775 | -95.2605
2.6811 | -95.2723
2.6847 | -95.3228
2.6883 | -95.4156
2.692 | -95.5447
2.6956 | -95.6934
2.6992 | -95.8418
2.7028 | -95.9489
2.7064 | -96.0063
2.71 | -95.983
2.7136 | -95.8826
2.7173 | -95.7455
2.7209 | -95.6089
2.7245 | -95.5163
2.7281 | -95.4909
2.7317 | -95.5445
2.7353 | -95.6669
2.7389 | -95.8139
2.7425 | -95.9482
2.7462 | -96.0537
2.7498 | -96.0941
2.7534 | -96.0269
2.757 | -95.8819
2.7606 | -95.6989
2.7642 | -95.5237
2.7678 | -95.4069
2.7715 | -95.3577
2.7751 | -95.3749
2.7787 | -95.4417
2.7823 | -95.546
2.7859 | -95.6813
2.7895 | -95.8108
2.7931 | -95.8892
2.7967 | -95.9093
2.8004 | -95.8592
2.804 | -95.7361
2.8076 | -95.5781
2.8112 | -95.4209
2.8148 | -95.2986
2.8184 | -95.241
2.822 | -95.2505
2.8257 | -95.3254
2.8293 | -95.4207
2.8329 | -95.5165
2.8365 | -95.6027
2.8401 | -95.6813
2.8437 | -95.7628
2.8473 | -95.8563
2.8509 | -95.959
2.8546 | -96.0743
2.8582 | -96.2065
2.8618 | -96.3469
2.8654 | -96.4615
2.869 | -96.5303
2.8726 | -96.5458
2.8762 | -96.5244
2.8799 | -96.4565
2.8835 | -96.3549
2.8871 | -96.2416
2.8907 | -96.1391
2.8943 | -96.0898
2.8979 | -96.0973
2.9015 | -96.1623
2.9051 | -96.2896
2.9088 | -96.4677
2.9124 | -96.6808
2.916 | -96.8872
2.9196 | -97.0206
2.9232 | -97.028
2.9268 | -96.8964
2.9304 | -96.6769
2.9341 | -96.444
2.9377 | -96.2597
2.9413 | -96.1556
2.9449 | -96.1407
2.9485 | -96.1996
2.9521 | -96.307
2.9557 | -96.4403
2.9594 | -96.5685
2.963 | -96.6623
2.9666 | -96.7136
2.9702 | -96.7301
2.9738 | -96.7187
2.9774 | -96.6679
2.981 | -96.6022
2.9846 | -96.5213
2.9883 | -96.4151
2.9919 | -96.3116
2.9955 | -96.2284
2.9991 | -96.193
3.0027 | -96.212
3.0063 | -96.2552
3.0099 | -96.2982
3.0135 | -96.344
3.0172 | -96.3837
3.0208 | -96.438
3.0244 | -96.519
3.028 | -96.622
3.0316 | -96.7554
3.0352 | -96.8886
3.0388 | -96.9806
3.0425 | -96.9872
3.0461 | -96.8897
3.0497 | -96.7124
3.0533 | -96.4997
3.0569 | -96.3033
3.0605 | -96.1626
3.0641 | -96.1203
3.0678 | -96.1779
3.0714 | -96.314
3.075 | -96.4783
3.0786 | -96.6189
3.0822 | -96.7131
3.0858 | -96.736
3.0894 | -96.7033
3.093 | -96.6481
3.0967 | -96.5975
3.1003 | -96.5843
3.1039 | -96.5962
3.1075 | -96.6115
3.1111 | -96.5996
3.1147 | -96.5402
3.1183 | -96.4562
3.122 | -96.3809
3.1256 | -96.3676
3.1292 | -96.3544
3.1328 | -96.3487
3.1364 | -96.3585
3.14 | -96.3825
3.1436 | -96.4129
3.1472 | -96.4494
3.1509 | -96.5017
3.1545 | -96.5676
3.1581 | -96.6544
3.1617 | -96.7634
3.1653 | -96.8897
3.1689 | -97.0034
3.1725 | -97.0709
3.1762 | -97.0691
3.1798 | -97.0068
3.1834 | -96.9055
3.187 | -96.79
3.1906 | -96.6674
3.1942 | -96.5542
3.1978 | -96.4558
3.2014 | -96.3842
3.2051 | -96.3192
3.2087 | -96.2796
3.2123 | -96.2588
3.2159 | -96.2606
3.2195 | -96.2854
3.2231 | -96.3488
3.2267 | -96.4582
3.2304 | -96.6093
3.234 | -96.7764
3.2376 | -96.9403
3.2412 | -97.0769
3.2448 | -97.1547
3.2484 | -97.1654
3.252 | -97.1134
3.2556 | -97.0273
3.2593 | -96.9389
3.2629 | -96.8536
3.2665 | -96.7946
3.2701 | -96.7745
3.2737 | -96.7967
3.2773 | -96.8442
3.2809 | -96.9117
3.2846 | -96.9927
3.2882 | -97.0768
3.2918 | -97.1501
3.2954 | -97.2047
3.299 | -97.2469
3.3026 | -97.275
3.3062 | -97.2937
3.3098 | -97.2992
3.3135 | -97.3013
3.3171 | -97.3175
3.3207 | -97.3249
3.3243 | -97.3037
3.3279 | -97.2536
3.3315 | -97.198
3.3351 | -97.1482
3.3388 | -97.1278
3.3424 | -97.1436
3.346 | -97.2003
3.3496 | -97.2884
3.3532 | -97.3881
3.3568 | -97.4908
3.3604 | -97.6048
3.364 | -97.7182
3.3677 | -97.8252
3.3713 | -97.9008
3.3749 | -97.9565
3.3785 | -97.9922
3.3821 | -98.0118
3.3857 | -98.0064
3.3893 | -97.9839
3.393 | -97.9602
3.3966 | -97.9384
3.4002 | -97.9098
3.4038 | -97.8806
3.4074 | -97.8579
3.411 | -97.8385
3.4146 | -97.8149
3.4182 | -97.7956
3.4219 | -97.7787
3.4255 | -97.7653
3.4291 | -97.7556
3.4327 | -97.7626
3.4363 | -97.791
3.4399 | -97.8282
3.4435 | -97.8603
3.4472 | -97.8679
3.4508 | -97.8437
3.4544 | -97.7708
3.458 | -97.668
3.4616 | -97.5413
3.4652 | -97.4095
3.4688 | -97.2942
3.4724 | -97.2057
3.4761 | -97.1619
3.4797 | -97.1604
3.4833 | -97.1963
3.4869 | -97.2475
3.4905 | -97.3046
3.4941 | -97.3727
3.4977 | -97.4437
3.5014 | -97.5134
3.505 | -97.5947
3.5086 | -97.6712
3.5122 | -97.7246
3.5158 | -97.7618
3.5194 | -97.7652
3.523 | -97.7403
3.5266 | -97.711
3.5303 | -97.6828
3.5339 | -97.6852
3.5375 | -97.7294
3.5411 | -97.7987
3.5447 | -97.8775
3.5483 | -97.959
3.5519 | -98.0347
3.5556 | -98.1066
3.5592 | -98.1851
3.5628 | -98.274
3.5664 | -98.3706
3.57 | -98.4678
3.5736 | -98.564
3.5772 | -98.632
3.5808 | -98.6456
3.5845 | -98.6048
3.5881 | -98.5106
3.5917 | -98.3636
3.5953 | -98.1914
3.5989 | -98.022
3.6025 | -97.8913
3.6061 | -97.818
3.6098 | -97.8038
3.6134 | -97.837
3.617 | -97.9152
3.6206 | -98.0103
3.6242 | -98.1115
3.6278 | -98.2169
3.6314 | -98.3213
3.6351 | -98.4145
3.6387 | -98.482
3.6423 | -98.5173
3.6459 | -98.4581
3.6495 | -98.4219
3.6531 | -98.3664
3.6567 | -98.2953
3.6603 | -98.2096
3.664 | -98.1499
3.6676 | -98.1239
3.6712 | -98.1226
3.6748 | -98.1439
3.6784 | -98.1849
3.682 | -98.2449
3.6856 | -98.318
3.6892 | -98.4021
3.6929 | -98.499
3.6965 | -98.5964
3.7001 | -98.679
3.7037 | -98.7449
3.7073 | -98.7722
3.7109 | -98.7658
3.7145 | -98.7407
3.7182 | -98.7116
3.7218 | -98.6839
3.7254 | -98.6357
3.729 | -98.5807
3.7326 | -98.5212
3.7362 | -98.4509
3.7398 | -98.3728
3.7435 | -98.3065
3.7471 | -98.2692
3.7507 | -98.2771
3.7543 | -98.3391
3.7579 | -98.4489
3.7615 | -98.5895
3.7651 | -98.7438
3.7687 | -98.8923
3.7724 | -99.0094
3.776 | -99.0856
3.7796 | -99.1242
3.7832 | -99.1268
3.7868 | -99.1086
3.7904 | -99.0657
3.794 | -98.9934
3.7977 | -98.8972
3.8013 | -98.7866
3.8049 | -98.6585
3.8085 | -98.5292
3.8121 | -98.4097
3.8157 | -98.3195
3.8193 | -98.2692
3.8229 | -98.2677
3.8266 | -98.3122
3.8302 | -98.4016
3.8338 | -98.5322
3.8374 | -98.6879
3.841 | -98.8479
3.8446 | -99.0103
3.8482 | -99.1541
3.8519 | -99.2754
3.8555 | -99.3801
3.8591 | -99.4661
3.8627 | -99.5244
3.8663 | -99.5633
3.8699 | -99.592
3.8735 | -99.6151
3.8771 | -99.6195
3.8808 | -99.603
3.8844 | -99.5681
3.888 | -99.5057
3.8916 | -99.4187
3.8952 | -99.3062
3.8988 | -99.1811
3.9024 | -99.0612
3.9061 | -98.9565
3.9097 | -98.8734
3.9133 | -98.8128
3.9169 | -98.7777
3.9205 | -98.7665
3.9241 | -98.7819
3.9277 | -98.8092
3.9313 | -98.8354
3.935 | -98.8695
3.9386 | -98.9081
3.9422 | -98.9566
3.9458 | -99.0056
3.9494 | -99.0546
3.953 | -99.1052
3.9566 | -99.1521
3.9603 | -99.1805
3.9639 | -99.1968
3.9675 | -99.2147
3.9711 | -99.2403
3.9747 | -99.2795
3.9783 | -99.3227
3.9819 | -99.3636
3.9855 | -99.4027
3.9892 | -99.4306
3.9928 | -99.4441
3.9964 | -99.431
4 | -99.4061
4.0036 | -99.3742
4.0072 | -99.3511
4.0108 | -99.3442
4.0145 | -99.348
4.0181 | -99.3596
4.0217 | -99.3763
4.0253 | -99.3676
4.0289 | -99.3269
4.0325 | -99.2583
4.0361 | -99.1737
4.0397 | -99.0886
4.0434 | -99.0151
4.047 | -98.9684
4.0506 | -98.97
4.0542 | -99.0235
4.0578 | -99.1179
4.0614 | -99.2266
4.065 | -99.3244
4.0687 | -99.3885
4.0723 | -99.4108
4.0759 | -99.3891
4.0795 | -99.3413
4.0831 | -99.2793
4.0867 | -99.225
4.0903 | -99.1807
4.0939 | -99.166
4.0976 | -99.1724
4.1012 | -99.1973
4.1048 | -99.2287
4.1084 | -99.2671
4.112 | -99.301
4.1156 | -99.3231
4.1192 | -99.3391
4.1229 | -99.3557
4.1265 | -99.3615
4.1301 | -99.3604
4.1337 | -99.3597
4.1373 | -99.3544
4.1409 | -99.3447
4.1445 | -99.3364
4.1481 | -99.3275
4.1518 | -99.3158
4.1554 | -99.2995
4.159 | -99.299
4.1626 | -99.3254
4.1662 | -99.3727
4.1698 | -99.4425
4.1734 | -99.5011
4.1771 | -99.5575
4.1807 | -99.6116
4.1843 | -99.6575
4.1879 | -99.687
4.1915 | -99.7085
4.1951 | -99.729
4.1987 | -99.7514
4.2023 | -99.7829
4.206 | -99.8162
4.2096 | -99.8441
4.2132 | -99.8565
4.2168 | -99.8546
4.2204 | -99.8267
4.224 | -99.7586
4.2276 | -99.6443
4.2313 | -99.5087
4.2349 | -99.3677
4.2385 | -99.2208
4.2421 | -99.1025
4.2457 | -99.0176
4.2493 | -98.9587
4.2529 | -98.9288
4.2565 | -98.9316
4.2602 | -98.9563
4.2638 | -98.995
4.2674 | -99.05
4.271 | -99.1137
4.2746 | -99.1842
4.2782 | -99.2639
4.2818 | -99.3522
4.2855 | -99.4328
4.2891 | -99.5041
4.2927 | -99.575
4.2963 | -99.6463
4.2999 | -99.7058
4.3035 | -99.7509
4.3071 | -99.7856
4.3107 | -99.7986
4.3144 | -99.793
4.318 | -99.7866
4.3216 | -99.7703
4.3252 | -99.7459
4.3288 | -99.7268
4.3324 | -99.7077
4.336 | -99.685
4.3397 | -99.6609
4.3433 | -99.6214
4.3469 | -99.5714
4.3505 | -99.5087
4.3541 | -99.4404
4.3577 | -99.38
4.3613 | -99.3167
4.365 | -99.2593
4.3686 | -99.2198
4.3722 | -99.1952
4.3758 | -99.1945
4.3794 | -99.2255
4.383 | -99.2896
4.3866 | -99.3691
4.3902 | -99.4639
4.3939 | -99.5818
4.3975 | -99.7062
4.4011 | -99.8356
4.4047 | -99.9642
4.4083 | -100.0813
4.4119 | -100.1793
4.4155 | -100.2335
4.4192 | -100.2554
4.4228 | -100.247
4.4264 | -100.2083
4.43 | -100.1559
4.4336 | -100.0922
4.4372 | -100.0256
4.4408 | -99.9694
4.4444 | -99.9213
4.4481 | -99.8807
4.4517 | -99.8461
4.4553 | -99.8251
4.4589 | -99.824
4.4625 | -99.8256
4.4661 | -99.8415
4.4697 | -99.8659
4.4734 | -99.8871
4.477 | -99.9029
4.4806 | -99.9132
4.4842 | -99.9209
4.4878 | -99.9303
4.4914 | -99.9453
4.495 | -99.9639
4.4986 | -99.9984
4.5023 | -100.0421
4.5059 | -100.0897
4.5095 | -100.1416
4.5131 | -100.1981
4.5167 | -100.2561
4.5203 | -100.3069
4.5239 | -100.3432
4.5276 | -100.3621
4.5312 | -100.369
4.5348 | -100.3574
4.5384 | -100.3327
4.542 | -100.3036
4.5456 | -100.2712
4.5492 | -100.2475
4.5528 | -100.2259
4.5565 | -100.1986
4.5601 | -100.1763
4.5637 | -100.1618
4.5673 | -100.1398
4.5709 | -100.1187
4.5745 | -100.0906
4.5781 | -100.0545
4.5818 | -100.0173
4.5854 | -99.9832
4.589 | -99.9595
4.5926 | -99.9531
4.5962 | -99.9615
4.5998 | -99.9921
4.6034 | -100.0478
4.607 | -100.1191
4.6107 | -100.1983
4.6143 | -100.2742
4.6179 | -100.3331
4.6215 | -100.3675
4.6251 | -100.3798
4.6287 | -100.3609
4.6323 | -100.3114
4.636 | -100.2411
4.6396 | -100.152
4.6432 | -100.0586
4.6468 | -99.9731
4.6504 | -99.9071
4.654 | -99.8756
4.6576 | -99.8843
4.6612 | -99.9277
4.6649 | -99.9997
4.6685 | -100.0916
4.6721 | -100.1996
4.6757 | -100.3049
4.6793 | -100.3922
4.6829 | -100.4549
4.6865 | -100.489
4.6902 | -100.4772
4.6938 | -100.4931
4.6974 | -100.5007
4.701 | -100.5046
4.7046 | -100.5152
4.7082 | -100.5325
4.7118 | -100.5532
4.7154 | -100.5808
4.7191 | -100.6121
4.7227 | -100.6452
4.7263 | -100.6765
4.7299 | -100.7053
4.7335 | -100.7368
4.7371 | -100.7715
4.7407 | -100.8012
4.7444 | -100.8162
4.748 | -100.8312
4.7516 | -100.8471
4.7552 | -100.8656
4.7588 | -100.8872
4.7624 | -100.9066
4.766 | -100.9401
4.7696 | -100.9748
4.7733 | -101.0075
4.7769 | -101.0395
4.7805 | -101.0551
4.7841 | -101.0556
4.7877 | -101.0314
4.7913 | -100.9819
4.7949 | -100.9053
4.7986 | -100.8062
4.8022 | -100.7085
4.8058 | -100.6252
4.8094 | -100.5557
4.813 | -100.509
4.8166 | -100.4856
4.8202 | -100.4903
4.8238 | -100.5163
4.8275 | -100.5565
4.8311 | -100.5994
4.8347 | -100.6565
4.8383 | -100.7189
4.8419 | -100.7662
4.8455 | -100.7997
4.8491 | -100.8197
4.8528 | -100.8226
4.8564 | -100.8142
4.86 | -100.7921
4.8636 | -100.7661
4.8672 | -100.737
4.8708 | -100.7213
4.8744 | -100.7103
4.878 | -100.704
4.8817 | -100.7056
4.8853 | -100.7222
4.8889 | -100.7527
4.8925 | -100.7938
4.8961 | -100.8389
4.8997 | -100.8901
4.9033 | -100.9517
4.907 | -101.0093
4.9106 | -101.0395
4.9142 | -101.0551
4.9178 | -101.0519
4.9214 | -101.0273
4.925 | -100.9807
4.9286 | -100.9212
4.9322 | -100.8622
4.9359 | -100.8207
4.9395 | -100.8015
4.9431 | -100.8072
4.9467 | -100.8235
4.9503 | -100.8644
4.9539 | -100.9121
4.9575 | -100.9623
4.9612 | -101.0091
4.9648 | -101.0439
4.9684 | -101.066
4.972 | -101.0696
4.9756 | -101.0512
4.9792 | -101.0117
4.9828 | -100.9734
4.9864 | -100.9357
4.9901 | -100.8865
4.9937 | -100.8531
4.9973 | -100.8333
5.0009 | -100.8222
5.0045 | -100.8337
5.0081 | -100.8524
5.0117 | -100.8785
5.0154 | -100.9056
5.019 | -100.9252
5.0226 | -100.9411
5.0262 | -100.945
5.0298 | -100.942
5.0334 | -100.9254
5.037 | -100.9042
5.0407 | -100.8895
5.0443 | -100.8849
5.0479 | -100.8893
5.0515 | -100.9063
5.0551 | -100.9409
5.0587 | -100.9825
5.0623 | -101.0325
5.0659 | -101.0931
5.0696 | -101.1489
5.0732 | -101.2019
5.0768 | -101.2562
5.0804 | -101.2984
5.084 | -101.3441
5.0876 | -101.3924
5.0912 | -101.4553
5.0949 | -101.5268
5.0985 | -101.5998
5.1021 | -101.6666
5.1057 | -101.7216
5.1093 | -101.7609
5.1129 | -101.7764
5.1165 | -101.7658
5.1201 | -101.7307
5.1238 | -101.6889
5.1274 | -101.639
5.131 | -101.5804
5.1346 | -101.5259
5.1382 | -101.484
5.1418 | -101.4475
5.1454 | -101.4199
5.1491 | -101.4086
5.1527 | -101.409
5.1563 | -101.4172
5.1599 | -101.4261
5.1635 | -101.4199
5.1671 | -101.4086
5.1707 | -101.3852
5.1743 | -101.3515
5.178 | -101.3109
5.1816 | -101.2765
5.1852 | -101.2467
5.1888 | -101.2161
5.1924 | -101.197
5.196 | -101.1969
5.1996 | -101.2089
5.2033 | -101.231
5.2069 | -101.263
5.2105 | -101.2878
5.2141 | -101.3114
5.2177 | -101.3343
5.2213 | -101.3531
5.2249 | -101.3595
5.2285 | -101.3556
5.2322 | -101.333
5.2358 | -101.2975
5.2394 | -101.2629
5.243 | -101.2241
5.2466 | -101.1735
5.2502 | -101.1264
5.2538 | -101.0848
5.2575 | -101.0504
5.2611 | -101.0196
5.2647 | -101.005
5.2683 | -101.0053
5.2719 | -101.0067
5.2755 | -101.019
5.2791 | -101.0426
5.2827 | -101.0787
5.2864 | -101.1164
5.29 | -101.153
5.2936 | -101.1751
5.2972 | -101.1735
5.3008 | -101.1543
5.3044 | -101.1206
5.308 | -101.0735
5.3117 | -101.0186
5.3153 | -100.969
5.3189 | -100.9245
5.3225 | -100.8768
5.3261 | -100.8438
5.3297 | -100.8273
5.3333 | -100.8144
5.3369 | -100.8057
5.3406 | -100.8149
5.3442 | -100.8357
5.3478 | -100.8606
5.3514 | -100.8928
5.355 | -100.9321
5.3586 | -100.9646
5.3622 | -100.9931
5.3659 | -101.0232
5.3695 | -101.052
5.3731 | -101.073
5.3767 | -101.0916
5.3803 | -101.1067
5.3839 | -101.107
5.3875 | -101.0945
5.3911 | -101.0781
5.3948 | -101.0617
5.3984 | -101.0429
5.402 | -101.036
5.4056 | -101.0353
5.4092 | -101.037
5.4128 | -101.0367
5.4164 | -101.0414
5.4201 | -101.0468
5.4237 | -101.0427
5.4273 | -101.0435
5.4309 | -101.0543
5.4345 | -101.0651
5.4381 | -101.0762
5.4417 | -101.0925
5.4453 | -101.1113
5.449 | -101.1212
5.4526 | -101.1367
5.4562 | -101.1565
5.4598 | -101.1678
5.4634 | -101.179
5.467 | -101.1942
5.4706 | -101.2088
5.4743 | -101.2194
5.4779 | -101.2239
5.4815 | -101.2225
5.4851 | -101.2158
5.4887 | -101.2187
5.4923 | -101.2284
5.4959 | -101.248
5.4995 | -101.2761
5.5032 | -101.3126
5.5068 | -101.3531
5.5104 | -101.3948
5.514 | -101.4343
5.5176 | -101.47
5.5212 | -101.5046
5.5248 | -101.5345
5.5285 | -101.5554
5.5321 | -101.568
5.5357 | -101.5609
5.5393 | -101.5436
5.5429 | -101.5236
5.5465 | -101.4938
5.5501 | -101.4535
5.5537 | -101.41
5.5574 | -101.3644
5.561 | -101.3238
5.5646 | -101.2929
5.5682 | -101.2602
5.5718 | -101.2328
5.5754 | -101.2172
5.579 | -101.2123
5.5827 | -101.2239
5.5863 | -101.2529
5.5899 | -101.3067
5.5935 | -101.3801
5.5971 | -101.4743
5.6007 | -101.5848
5.6043 | -101.6923
5.6079 | -101.7865
5.6116 | -101.8689
5.6152 | -101.9375
5.6188 | -101.9881
5.6224 | -102.0184
5.626 | -102.0301
5.6296 | -102.0392
5.6332 | -102.0523
5.6369 | -102.0699
5.6405 | -102.0891
5.6441 | -102.1033
5.6477 | -102.1161
5.6513 | -102.1284
5.6549 | -102.1236
5.6585 | -102.1016
5.6621 | -102.0658
5.6658 | -102.0205
5.6694 | -101.9728
5.673 | -101.9271
5.6766 | -101.8875
5.6802 | -101.8522
5.6838 | -101.8206
5.6874 | -101.8012
5.6911 | -101.7925
5.6947 | -101.7876
5.6983 | -101.7921
5.7019 | -101.8018
5.7055 | -101.8162
5.7091 | -101.8403
5.7127 | -101.8633
5.7164 | -101.8846
5.72 | -101.909
5.7236 | -101.9348
5.7272 | -101.9634
5.7308 | -101.9848
5.7344 | -102.0059
5.738 | -102.0184
5.7416 | -102.0287
5.7453 | -102.0339
5.7489 | -102.0315
5.7525 | -102.0209
5.7561 | -102.0022
5.7597 | -101.9698
5.7633 | -101.928
5.7669 | -101.8936
5.7706 | -101.8687
5.7742 | -101.8455
5.7778 | -101.836
5.7814 | -101.8398
5.785 | -101.8568
5.7886 | -101.8956
5.7922 | -101.9528
5.7958 | -102.0172
5.7995 | -102.0836
5.8031 | -102.1521
5.8067 | -102.2153
5.8103 | -102.2678
5.8139 | -102.3112
5.8175 | -102.3433
5.8211 | -102.3604
5.8248 | -102.3624
5.8284 | -102.3533
5.832 | -102.3372
5.8356 | -102.3129
5.8392 | -102.282
5.8428 | -102.247
5.8464 | -102.2219
5.85 | -102.2092
5.8537 | -102.193
5.8573 | -102.1651
5.8609 | -102.1411
5.8645 | -102.1184
5.8681 | -102.0849
5.8717 | -102.0471
5.8753 | -102.0033
5.879 | -101.9501
5.8826 | -101.889
5.8862 | -101.8378
5.8898 | -101.7966
5.8934 | -101.7577
5.897 | -101.7265
5.9006 | -101.7013
5.9042 | -101.675
5.9079 | -101.6516
5.9115 | -101.6294
5.9151 | -101.6077
5.9187 | -101.5926
5.9223 | -101.579
5.9259 | -101.5681
5.9295 | -101.5647
5.9332 | -101.5649
5.9368 | -101.5581
5.9404 | -101.5512
5.944 | -101.5524
5.9476 | -101.5508
5.9512 | -101.5357
5.9548 | -101.5217
5.9584 | -101.5126
5.9621 | -101.4986
5.9657 | -101.4854
5.9693 | -101.4822
5.9729 | -101.4853
5.9765 | -101.4862
5.9801 | -101.4928
5.9837 | -101.5074
5.9874 | -101.5273
5.991 | -101.5545
5.9946 | -101.5893
5.9982 | -101.6225
6.0018 | -101.6557
6.0054 | -101.6889
6.009 | -101.714
6.0126 | -101.7317
6.0163 | -101.7434
6.0199 | -101.754
6.0235 | -101.764
6.0271 | -101.7718
6.0307 | -101.7764
6.0343 | -101.7836
6.0379 | -101.7918
6.0416 | -101.8037
6.0452 | -101.8199
6.0488 | -101.8426
6.0524 | -101.8781
6.056 | -101.9263
6.0596 | -101.9798
6.0632 | -102.0443
6.0668 | -102.1201
6.0705 | -102.1972
6.0741 | -102.2698
6.0777 | -102.3401
6.0813 | -102.3943
6.0849 | -102.4316
6.0885 | -102.4598
6.0921 | -102.4704
6.0958 | -102.4615
6.0994 | -102.4345
6.103 | -102.3905
6.1066 | -102.3231
6.1102 | -102.2373
6.1138 | -102.1417
6.1174 | -102.0382
6.121 | -101.9311
6.1247 | -101.8202
6.1283 | -101.7209
6.1319 | -101.6306
6.1355 | -101.5501
6.1391 | -101.4887
6.1427 | -101.4488
6.1463 | -101.4285
6.15 | -101.4368
6.1536 | -101.4737
6.1572 | -101.5317
6.1608 | -101.6123
6.1644 | -101.7095
6.168 | -101.8113
6.1716 | -101.9213
6.1752 | -102.028
6.1789 | -102.124
6.1825 | -102.2153
6.1861 | -102.2982
6.1897 | -102.3642
6.1933 | -102.4145
6.1969 | -102.4551
6.2005 | -102.4781
6.2042 | -102.4861
6.2078 | -102.4835
6.2114 | -102.4611
6.215 | -102.4322
6.2186 | -102.3887
6.2222 | -102.3447
6.2258 | -102.3028
6.2294 | -102.2548
6.2331 | -102.2147
6.2367 | -102.1798
6.2403 | -102.144
6.2439 | -102.1106
6.2475 | -102.0854
6.2511 | -102.0595
6.2547 | -102.0447
6.2584 | -102.0295
6.262 | -102.0223
6.2656 | -102.0234
6.2692 | -102.03
6.2728 | -102.0455
6.2764 | -102.0668
6.28 | -102.0943
6.2836 | -102.1304
6.2873 | -102.1657
6.2909 | -102.1942
6.2945 | -102.2195
6.2981 | -102.2433
6.3017 | -102.2607
6.3053 | -102.2785
6.3089 | -102.2872
6.3126 | -102.2977
6.3162 | -102.3051
6.3198 | -102.3151
6.3234 | -102.3341
6.327 | -102.357
6.3306 | -102.3853
6.3342 | -102.4201
6.3379 | -102.4676
6.3415 | -102.5243
6.3451 | -102.5905
6.3487 | -102.6683
6.3523 | -102.7443
6.3559 | -102.8137
6.3595 | -102.8756
6.3631 | -102.9249
6.3668 | -102.9619
6.3704 | -102.9737
6.374 | -102.9667
6.3776 | -102.9457
6.3812 | -102.9062
6.3848 | -102.8473
6.3884 | -102.7806
6.3921 | -102.7005
6.3957 | -102.6177
6.3993 | -102.5442
6.4029 | -102.4738
6.4065 | -102.4077
6.4101 | -102.3531
6.4137 | -102.3122
6.4173 | -102.2825
6.421 | -102.2673
6.4246 | -102.2692
6.4282 | -102.2907
6.4318 | -102.3268
6.4354 | -102.3691
6.439 | -102.4252
6.4426 | -102.4921
6.4463 | -102.554
6.4499 | -102.6182
6.4535 | -102.6787
6.4571 | -102.7307
6.4607 | -102.7705
6.4643 | -102.7988
6.4679 | -102.8233
6.4715 | -102.8463
6.4752 | -102.8694
6.4788 | -102.8974
6.4824 | -102.9335
6.486 | -102.9671
6.4896 | -103.0053
6.4932 | -103.0471
6.4968 | -103.0876
6.5005 | -103.1253
6.5041 | -103.1625
6.5077 | -103.1952
6.5113 | -103.2221
6.5149 | -103.2388
6.5185 | -103.2481
6.5221 | -103.2434
6.5257 | -103.2292
6.5294 | -103.2093
6.533 | -103.1827
6.5366 | -103.1476
6.5402 | -103.1062
6.5438 | -103.0693
6.5474 | -103.0337
6.551 | -102.9978
6.5547 | -102.9712
6.5583 | -102.9467
6.5619 | -102.9229
6.5655 | -102.8995
6.5691 | -102.8684
6.5727 | -102.8343
6.5763 | -102.7969
6.5799 | -102.7558
6.5836 | -102.7104
6.5872 | -102.6605
6.5908 | -102.6063
6.5944 | -102.5468
6.598 | -102.4917
6.6016 | -102.4445
6.6052 | -102.4031
6.6089 | -102.3705
6.6125 | -102.3566
6.6161 | -102.3583
6.6197 | -102.372
6.6233 | -102.4011
6.6269 | -102.4409
6.6305 | -102.4874
6.6341 | -102.5429
6.6378 | -102.6003
6.6414 | -102.6485
6.645 | -102.6861
6.6486 | -102.7163
6.6522 | -102.7283
6.6558 | -102.7286
6.6594 | -102.7138
6.6631 | -102.6783
6.6667 | -102.6387
6.6703 | -102.6002
6.6739 | -102.5573
6.6775 | -102.5111
6.6811 | -102.4665
6.6847 | -102.4279
6.6883 | -102.3943
6.692 | -102.3685
6.6956 | -102.3508
6.6992 | -102.3367
6.7028 | -102.3308
6.7064 | -102.3355
6.71 | -102.3432
6.7136 | -102.3591
6.7173 | -102.3874
6.7209 | -102.4263
6.7245 | -102.4742
6.7281 | -102.5279
6.7317 | -102.5854
6.7353 | -102.65
6.7389 | -102.7138
6.7425 | -102.7659
6.7462 | -102.8038
6.7498 | -102.8287
6.7534 | -102.8428
6.757 | -102.851
6.7606 | -102.8561
6.7642 | -102.8426
6.7678 | -102.8153
6.7715 | -102.7919
6.7751 | -102.7725
6.7787 | -102.7559
6.7823 | -102.7453
6.7859 | -102.7398
6.7895 | -102.7376
6.7931 | -102.7475
6.7967 | -102.7681
6.8004 | -102.8004
6.804 | -102.8427
6.8076 | -102.8894
6.8112 | -102.9454
6.8148 | -102.9942
6.8184 | -103.0364
6.822 | -103.0681
6.8257 | -103.0883
6.8293 | -103.0918
6.8329 | -103.082
6.8365 | -103.0576
6.8401 | -103.0211
6.8437 | -102.981
6.8473 | -102.9353
6.8509 | -102.883
6.8546 | -102.8291
6.8582 | -102.7725
6.8618 | -102.7179
6.8654 | -102.6767
6.869 | -102.6399
6.8726 | -102.6095
6.8762 | -102.5881
6.8799 | -102.575
6.8835 | -102.5705
6.8871 | -102.5733
6.8907 | -102.5828
6.8943 | -102.6044
6.8979 | -102.6295
6.9015 | -102.6573
6.9051 | -102.6945
6.9088 | -102.7346
6.9124 | -102.7831
6.916 | -102.8339
6.9196 | -102.8885
6.9232 | -102.9414
6.9268 | -102.9953
6.9304 | -103.052
6.9341 | -103.1124
6.9377 | -103.1636
6.9413 | -103.2139
6.9449 | -103.2705
6.9485 | -103.3195
6.9521 | -103.3608
6.9557 | -103.401
6.9593 | -103.4313
6.963 | -103.4479
6.9666 | -103.4562
6.9702 | -103.4567
6.9738 | -103.4525
6.9774 | -103.4352
6.981 | -103.4084
6.9846 | -103.3795
6.9883 | -103.3442
6.9919 | -103.3005
6.9955 | -103.2588
6.9991 | -103.2202
7.0027 | -103.1895
7.0063 | -103.1633
7.0099 | -103.1427
7.0136 | -103.1323
7.0172 | -103.1253
7.0208 | -103.128
7.0244 | -103.1396
7.028 | -103.1534
7.0316 | -103.1691
7.0352 | -103.1879
7.0388 | -103.203
7.0425 | -103.2242
7.0461 | -103.2442
7.0497 | -103.253
7.0533 | -103.2535
7.0569 | -103.2573
7.0605 | -103.2486
7.0641 | -103.2341
7.0678 | -103.2242
7.0714 | -103.2129
7.075 | -103.1925
7.0786 | -103.1709
7.0822 | -103.1585
7.0858 | -103.1526
7.0894 | -103.1489
7.093 | -103.1477
7.0967 | -103.1516
7.1003 | -103.1559
7.1039 | -103.1614
7.1075 | -103.1641
7.1111 | -103.1656
7.1147 | -103.1642
7.1183 | -103.1559
7.122 | -103.1441
7.1256 | -103.1245
7.1292 | -103.1003
7.1328 | -103.0787
7.1364 | -103.0521
7.14 | -103.0272
7.1436 | -103.0087
7.1472 | -102.9862
7.1509 | -102.9678
7.1545 | -102.9589
7.1581 | -102.9607
7.1617 | -102.9551
7.1653 | -102.9591
7.1689 | -102.9668
7.1725 | -102.9757
7.1762 | -102.9832
7.1798 | -102.9991
7.1834 | -103.0196
7.187 | -103.0414
7.1906 | -103.0611
7.1942 | -103.0832
7.1978 | -103.1087
7.2014 | -103.1351
7.2051 | -103.1611
7.2087 | -103.189
7.2123 | -103.2205
7.2159 | -103.2526
7.2195 | -103.2813
7.2231 | -103.3139
7.2267 | -103.3422
7.2304 | -103.3662
7.234 | -103.3891
7.2376 | -103.4029
7.2412 | -103.4119
7.2448 | -103.4137
7.2484 | -103.4105
7.252 | -103.4036
7.2556 | -103.391
7.2593 | -103.3764
7.2629 | -103.3656
7.2665 | -103.3614
7.2701 | -103.3562
7.2737 | -103.3518
7.2773 | -103.3481
7.2809 | -103.3539
7.2846 | -103.3589
7.2882 | -103.3616
7.2918 | -103.363
7.2954 | -103.3628
7.299 | -103.3621
7.3026 | -103.3578
7.3062 | -103.3486
7.3098 | -103.3336
7.3135 | -103.3147
7.3171 | -103.2969
7.3207 | -103.2828
7.3243 | -103.2736
7.3279 | -103.2639
7.3315 | -103.2521
7.3351 | -103.248
7.3388 | -103.248
7.3424 | -103.2522
7.346 | -103.2617
7.3496 | -103.2726
7.3532 | -103.2842
7.3568 | -103.2931
7.3604 | -103.3024
7.364 | -103.3106
7.3677 | -103.315
7.3713 | -103.3115
7.3749 | -103.306
7.3785 | -103.3003
7.3821 | -103.2936
7.3857 | -103.2793
7.3893 | -103.2561
7.393 | -103.2374
7.3966 | -103.2199
7.4002 | -103.1999
7.4038 | -103.1822
7.4074 | -103.1654
7.411 | -103.1568
7.4146 | -103.151
7.4182 | -103.1506
7.4219 | -103.1509
7.4255 | -103.1489
7.4291 | -103.1474
7.4327 | -103.1475
7.4363 | -103.1512
7.4399 | -103.1595
7.4435 | -103.1654
7.4472 | -103.1683
7.4508 | -103.1691
7.4544 | -103.1627
7.458 | -103.1556
7.4616 | -103.147
7.4652 | -103.1298
7.4688 | -103.1144
7.4724 | -103.0869
7.4761 | -103.0524
7.4797 | -103.0155
7.4833 | -102.9777
7.4869 | -102.9381
7.4905 | -102.897
7.4941 | -102.854
7.4977 | -102.8185
7.5014 | -102.7884
7.505 | -102.7542
7.5086 | -102.7285
7.5122 | -102.7127
7.5158 | -102.7009
7.5194 | -102.6964
7.523 | -102.703
7.5266 | -102.7212
7.5303 | -102.7489
7.5339 | -102.7879
7.5375 | -102.8366
7.5411 | -102.8914
7.5447 | -102.9537
7.5483 | -103.0165
7.5519 | -103.0862
7.5556 | -103.1608
7.5592 | -103.2269
7.5628 | -103.2907
7.5664 | -103.356
7.57 | -103.4062
7.5736 | -103.4523
7.5772 | -103.4978
7.5808 | -103.5296
7.5845 | -103.5626
7.5881 | -103.5882
7.5917 | -103.6076
7.5953 | -103.6222
7.5989 | -103.631
7.6025 | -103.6344
7.6061 | -103.6336
7.6098 | -103.6276
7.6134 | -103.6202
7.617 | -103.6192
7.6206 | -103.614
7.6242 | -103.6057
7.6278 | -103.5998
7.6314 | -103.5908
7.635 | -103.5774
7.6387 | -103.5613
7.6423 | -103.5402
7.6459 | -103.5259
7.6495 | -103.5185
7.6531 | -103.5128
7.6567 | -103.5129
7.6603 | -103.5283
7.664 | -103.5474
7.6676 | -103.5765
7.6712 | -103.6165
7.6748 | -103.663
7.6784 | -103.7065
7.682 | -103.7575
7.6856 | -103.8077
7.6893 | -103.8554
7.6929 | -103.897
7.6965 | -103.9326
7.7001 | -103.9641
7.7037 | -103.9872
7.7073 | -104.0047
7.7109 | -104.0136
7.7145 | -104.0128
7.7182 | -104.0057
7.7218 | -103.986
7.7254 | -103.9594
7.729 | -103.9289
7.7326 | -103.8963
7.7362 | -103.8629
7.7398 | -103.8318
7.7435 | -103.8034
7.7471 | -103.7671
7.7507 | -103.7312
7.7543 | -103.6982
7.7579 | -103.6616
7.7615 | -103.6252
7.7651 | -103.5883
7.7687 | -103.552
7.7724 | -103.5191
7.776 | -103.4911
7.7796 | -103.462
7.7832 | -103.439
7.7868 | -103.4202
7.7904 | -103.4099
7.794 | -103.4041
7.7977 | -103.4
7.8013 | -103.4037
7.8049 | -103.4172
7.8085 | -103.4308
7.8121 | -103.4438
7.8157 | -103.4613
7.8193 | -103.486
7.8229 | -103.506
7.8266 | -103.5295
7.8302 | -103.5542
7.8338 | -103.5769
7.8374 | -103.5986
7.841 | -103.6204
7.8446 | -103.6383
7.8482 | -103.6554
7.8519 | -103.6741
7.8555 | -103.6928
7.8591 | -103.7116
7.8627 | -103.7394
7.8663 | -103.7641
7.8699 | -103.7816
7.8735 | -103.796
7.8771 | -103.8131
7.8808 | -103.8341
7.8844 | -103.855
7.888 | -103.8701
7.8916 | -103.8851
7.8952 | -103.9057
7.8988 | -103.923
7.9024 | -103.9366
7.9061 | -103.9509
7.9097 | -103.9642
7.9133 | -103.9777
7.9169 | -103.9864
7.9205 | -103.9908
7.9241 | -103.995
7.9277 | -104.0001
7.9313 | -104.0034
7.935 | -104.0023
7.9386 | -104.0007
7.9422 | -103.9984
7.9458 | -104.0015
7.9494 | -104.0041
7.953 | -104.004
7.9566 | -104.0053
7.9603 | -104.0038
7.9639 | -104.0047
7.9675 | -104.008
7.9711 | -104.008
7.9747 | -104.0069
7.9783 | -104.0042
7.9819 | -103.9913
7.9855 | -103.981
7.9892 | -103.9711
7.9928 | -103.9582
7.9964 | -103.9402
8 | -103.9211
8.0036 | -103.9068
8.0072 | -103.8928
8.0108 | -103.8831
8.0145 | -103.8708
8.0181 | -103.8644
8.0217 | -103.8622
8.0253 | -103.8614
8.0289 | -103.8653
8.0325 | -103.8729
8.0361 | -103.8844
8.0397 | -103.903
8.0434 | -103.9201
8.047 | -103.9303
8.0506 | -103.9409
8.0542 | -103.9553
8.0578 | -103.9669
8.0614 | -103.9741
8.065 | -103.9756
8.0687 | -103.9751
8.0723 | -103.9764
8.0759 | -103.9737
8.0795 | -103.9712
8.0831 | -103.9673
8.0867 | -103.9618
8.0903 | -103.9562
8.0939 | -103.9461
8.0976 | -103.9328
8.1012 | -103.9203
8.1048 | -103.9003
8.1084 | -103.8844
8.112 | -103.8672
8.1156 | -103.8463
8.1192 | -103.8257
8.1229 | -103.8083
8.1265 | -103.7953
8.1301 | -103.7805
8.1337 | -103.7657
8.1373 | -103.7479
8.1409 | -103.7334
8.1445 | -103.7243
8.1481 | -103.7143
8.1518 | -103.7037
8.1554 | -103.6965
8.159 | -103.6926
8.1626 | -103.6957
8.1662 | -103.7019
8.1698 | -103.7118
8.1734 | -103.7231
8.1771 | -103.7407
8.1807 | -103.7664
8.1843 | -103.7958
8.1879 | -103.8274
8.1915 | -103.8576
8.1951 | -103.8914
8.1987 | -103.9211
8.2023 | -103.9473
8.206 | -103.9784
8.2096 | -104.0025
8.2132 | -104.0265
8.2168 | -104.0496
8.2204 | -104.0705
8.224 | -104.0876
8.2276 | -104.0972
8.2313 | -104.1031
8.2349 | -104.1072
8.2385 | -104.1057
8.2421 | -104.0984
8.2457 | -104.0888
8.2493 | -104.0733
8.2529 | -104.0601
8.2565 | -104.0493
8.2602 | -104.0324
8.2638 | -104.0105
8.2674 | -103.994
8.271 | -103.9825
8.2746 | -103.9708
8.2782 | -103.9593
8.2818 | -103.9505
8.2855 | -103.9465
8.2891 | -103.9424
8.2927 | -103.9407
8.2963 | -103.9422
8.2999 | -103.9459
8.3035 | -103.9532
8.3071 | -103.9655
8.3108 | -103.9797
8.3144 | -103.9955
8.318 | -104.0191
8.3216 | -104.0486
8.3252 | -104.0794
8.3288 | -104.1169
8.3324 | -104.1551
8.336 | -104.1913
8.3397 | -104.2163
8.3433 | -104.2435
8.3469 | -104.2686
8.3505 | -104.2884
8.3541 | -104.3053
8.3577 | -104.319
8.3613 | -104.3273
8.365 | -104.3342
8.3686 | -104.3408
8.3722 | -104.3443
8.3758 | -104.3412
8.3794 | -104.3314
8.383 | -104.3202
8.3866 | -104.3067
8.3902 | -104.2842
8.3939 | -104.2555
8.3975 | -104.2244
8.4011 | -104.1933
8.4047 | -104.1532
8.4083 | -104.1092
8.4119 | -104.0638
8.4155 | -104.0182
8.4192 | -103.9771
8.4228 | -103.9368
8.4264 | -103.8992
8.43 | -103.8686
8.4336 | -103.8474
8.4372 | -103.8273
8.4408 | -103.8104
8.4444 | -103.7989
8.4481 | -103.7889
8.4517 | -103.7883
8.4553 | -103.7922
8.4589 | -103.7992
8.4625 | -103.8124
8.4661 | -103.8312
8.4697 | -103.8527
8.4734 | -103.8715
8.477 | -103.8942
8.4806 | -103.9208
8.4842 | -103.9475
8.4878 | -103.9728
8.4914 | -103.9915
8.495 | -104.0075
8.4986 | -104.0218
8.5023 | -104.0294
8.5059 | -104.0325
8.5095 | -104.0358
8.5131 | -104.0339
8.5167 | -104.0246
8.5203 | -104.0124
8.5239 | -103.9956
8.5276 | -103.9735
8.5312 | -103.9538
8.5348 | -103.9379
8.5384 | -103.92
8.542 | -103.9037
8.5456 | -103.8833
8.5492 | -103.8685
8.5528 | -103.8558
8.5565 | -103.8445
8.5601 | -103.8355
8.5637 | -103.8352
8.5673 | -103.8387
8.5709 | -103.8427
8.5745 | -103.8548
8.5781 | -103.8735
8.5818 | -103.8964
8.5854 | -103.9235
8.589 | -103.9562
8.5926 | -103.9922
8.5962 | -104.0318
8.5998 | -104.0744
8.6034 | -104.1147
8.607 | -104.1554
8.6107 | -104.1937
8.6143 | -104.2326
8.6179 | -104.2732
8.6215 | -104.309
8.6251 | -104.3392
8.6287 | -104.3631
8.6323 | -104.3817
8.636 | -104.3914
8.6396 | -104.3984
8.6432 | -104.4076
8.6468 | -104.4081
8.6504 | -104.4014
8.654 | -104.3867
8.6576 | -104.3698
8.6612 | -104.3432
8.6649 | -104.3188
8.6685 | -104.2963
8.6721 | -104.2813
8.6757 | -104.2672
8.6793 | -104.2548
8.6829 | -104.2398
8.6865 | -104.2302
8.6902 | -104.2228
8.6938 | -104.2216
8.6974 | -104.2272
8.701 | -104.2339
8.7046 | -104.2351
8.7082 | -104.2366
8.7118 | -104.2446
8.7154 | -104.2522
8.7191 | -104.2569
8.7227 | -104.2629
8.7263 | -104.2711
8.7299 | -104.2781
8.7335 | -104.2807
8.7371 | -104.2796
8.7407 | -104.278
8.7444 | -104.2706
8.748 | -104.2633
8.7516 | -104.2521
8.7552 | -104.2382
8.7588 | -104.2241
8.7624 | -104.2118
8.766 | -104.2002
8.7696 | -104.1863
8.7733 | -104.1747
8.7769 | -104.1708
8.7805 | -104.1714
8.7841 | -104.167
8.7877 | -104.175
8.7913 | -104.1879
8.7949 | -104.2033
8.7986 | -104.2196
8.8022 | -104.2391
8.8058 | -104.2661
8.8094 | -104.2963
8.813 | -104.3232
8.8166 | -104.348
8.8202 | -104.368
8.8238 | -104.3847
8.8275 | -104.3928
8.8311 | -104.4012
8.8347 | -104.4084
8.8383 | -104.4074
8.8419 | -104.3984
8.8455 | -104.3864
8.8491 | -104.3727
8.8528 | -104.3577
8.8564 | -104.3399
8.86 | -104.3173
8.8636 | -104.2972
8.8672 | -104.2832
8.8708 | -104.2651
8.8744 | -104.2447
8.878 | -104.2205
8.8817 | -104.1991
8.8853 | -104.1766
8.8889 | -104.151
8.8925 | -104.1224
8.8961 | -104.0954
8.8997 | -104.0707
8.9033 | -104.0466
8.907 | -104.0223
8.9106 | -104.0024
8.9142 | -103.9861
8.9178 | -103.9645
8.9214 | -103.9453
8.925 | -103.9315
8.9286 | -103.9261
8.9322 | -103.9229
8.9359 | -103.9249
8.9395 | -103.9323
8.9431 | -103.9434
8.9467 | -103.9569
8.9503 | -103.9734
8.9539 | -103.9934
8.9575 | -104.0185
8.9612 | -104.0446
8.9648 | -104.0673
8.9684 | -104.0916
8.972 | -104.1224
8.9756 | -104.156
8.9792 | -104.1867
8.9828 | -104.2144
8.9864 | -104.2434
8.9901 | -104.2763
8.9937 | -104.3105
8.9973 | -104.3401
9.0009 | -104.3659
9.0045 | -104.394
9.0081 | -104.4216
9.0117 | -104.4533
9.0154 | -104.4824
9.019 | -104.5027
9.0226 | -104.5168
9.0262 | -104.5309
9.0298 | -104.5376
9.0334 | -104.5406
9.037 | -104.5397
9.0406 | -104.5288
9.0443 | -104.51
9.0479 | -104.4933
9.0515 | -104.4712
9.0551 | -104.4435
9.0587 | -104.4141
9.0623 | -104.3829
9.0659 | -104.3496
9.0696 | -104.3152
9.0732 | -104.2898
9.0768 | -104.2695
9.0804 | -104.2551
9.084 | -104.2455
9.0876 | -104.2394
9.0912 | -104.2393
9.0949 | -104.2435
9.0985 | -104.2569
9.1021 | -104.2728
9.1057 | -104.2904
9.1093 | -104.3113
9.1129 | -104.3353
9.1165 | -104.3625
9.1201 | -104.3864
9.1238 | -104.4045
9.1274 | -104.4277
9.131 | -104.449
9.1346 | -104.4673
9.1382 | -104.4827
9.1418 | -104.4966
9.1454 | -104.5054
9.1491 | -104.5143
9.1527 | -104.5229
9.1563 | -104.5254
9.1599 | -104.5241
9.1635 | -104.5232
9.1671 | -104.521
9.1707 | -104.5176
9.1743 | -104.5138
9.178 | -104.5054
9.1816 | -104.4905
9.1852 | -104.4768
9.1888 | -104.4634
9.1924 | -104.4438
9.196 | -104.4213
9.1996 | -104.4053
9.2033 | -104.3867
9.2069 | -104.3613
9.2105 | -104.3397
9.2141 | -104.3214
9.2177 | -104.3051
9.2213 | -104.2918
9.2249 | -104.2791
9.2285 | -104.2714
9.2322 | -104.2649
9.2358 | -104.2631
9.2394 | -104.2684
9.243 | -104.2777
9.2466 | -104.294
9.2502 | -104.3151
9.2538 | -104.3404
9.2575 | -104.3679
9.2611 | -104.3932
9.2647 | -104.4194
9.2683 | -104.4474
9.2719 | -104.4709
9.2755 | -104.4927
9.2791 | -104.5154
9.2827 | -104.5333
9.2864 | -104.5485
9.29 | -104.567
9.2936 | -104.5842
9.2972 | -104.5984
9.3008 | -104.6065
9.3044 | -104.6133
9.308 | -104.6182
9.3117 | -104.6299
9.3153 | -104.6438
9.3189 | -104.6517
9.3225 | -104.6591
9.3261 | -104.6702
9.3297 | -104.6774
9.3333 | -104.6795
9.3369 | -104.6841
9.3406 | -104.6862
9.3442 | -104.6865
9.3478 | -104.6905
9.3514 | -104.6952
9.355 | -104.6987
9.3586 | -104.703
9.3622 | -104.7065
9.3659 | -104.7097
9.3695 | -104.706
9.3731 | -104.6961
9.3767 | -104.6894
9.3803 | -104.6695
9.3839 | -104.6495
9.3875 | -104.6223
9.3911 | -104.5996
9.3948 | -104.5743
9.3984 | -104.544
9.402 | -104.5135
9.4056 | -104.4798
9.4092 | -104.4473
9.4128 | -104.4151
9.4164 | -104.3811
9.4201 | -104.3484
9.4237 | -104.3189
9.4273 | -104.2926
9.4309 | -104.2698
9.4345 | -104.2541
9.4381 | -104.2456
9.4417 | -104.245
9.4453 | -104.254
9.449 | -104.2622
9.4526 | -104.2744
9.4562 | -104.297
9.4598 | -104.3255
9.4634 | -104.3584
9.467 | -104.3979
9.4706 | -104.4414
9.4743 | -104.4822
9.4779 | -104.5276
9.4815 | -104.5701
9.4851 | -104.6092
9.4887 | -104.6472
9.4923 | -104.6818
9.4959 | -104.7126
9.4995 | -104.7348
9.5032 | -104.7497
9.5068 | -104.7655
9.5104 | -104.7764
9.514 | -104.7758
9.5176 | -104.7727
9.5212 | -104.7643
9.5248 | -104.7589
9.5285 | -104.7529
9.5321 | -104.7361
9.5357 | -104.7197
9.5393 | -104.7052
9.5429 | -104.6918
9.5465 | -104.6818
9.5501 | -104.6708
9.5537 | -104.6656
9.5574 | -104.6633
9.561 | -104.6656
9.5646 | -104.6745
9.5682 | -104.6831
9.5718 | -104.6924
9.5754 | -104.7089
9.579 | -104.7276
9.5827 | -104.7426
9.5863 | -104.7585
9.5899 | -104.7752
9.5935 | -104.7905
9.5971 | -104.8058
9.6007 | -104.8153
9.6043 | -104.8206
9.6079 | -104.8243
9.6116 | -104.822
9.6152 | -104.8183
9.6188 | -104.809
9.6224 | -104.7997
9.626 | -104.7901
9.6296 | -104.7784
9.6332 | -104.7676
9.6369 | -104.7608
9.6405 | -104.7609
9.6441 | -104.7649
9.6477 | -104.7656
9.6513 | -104.7651
9.6549 | -104.7695
9.6585 | -104.7834
9.6622 | -104.801
9.6658 | -104.8189
9.6694 | -104.8388
9.673 | -104.8632
9.6766 | -104.8848
9.6802 | -104.908
9.6838 | -104.929
9.6874 | -104.9484
9.6911 | -104.9693
9.6947 | -104.9837
9.6983 | -104.9954
9.7019 | -105.0034
9.7055 | -104.9999
9.7091 | -104.9906
9.7127 | -104.978
9.7164 | -104.961
9.72 | -104.944
9.7236 | -104.9269
9.7272 | -104.9027
9.7308 | -104.8788
9.7344 | -104.8564
9.738 | -104.828
9.7416 | -104.7968
9.7453 | -104.7709
9.7489 | -104.7486
9.7525 | -104.7277
9.7561 | -104.7114
9.7597 | -104.6997
9.7633 | -104.6875
9.7669 | -104.6758
9.7706 | -104.6701
9.7742 | -104.6664
9.7778 | -104.6584
9.7814 | -104.6505
9.785 | -104.6482
9.7886 | -104.6515
9.7922 | -104.6566
9.7958 | -104.6631
9.7995 | -104.6654
9.8031 | -104.6737
9.8067 | -104.6838
9.8103 | -104.6959
9.8139 | -104.7091
9.8175 | -104.7241
9.8211 | -104.7381
9.8248 | -104.75
9.8284 | -104.7605
9.832 | -104.7731
9.8356 | -104.7851
9.8392 | -104.797
9.8428 | -104.8066
9.8464 | -104.8125
9.85 | -104.8183
9.8537 | -104.8189
9.8573 | -104.8132
9.8609 | -104.8041
9.8645 | -104.7907
9.8681 | -104.7819
9.8717 | -104.7747
9.8753 | -104.7656
9.879 | -104.7569
9.8826 | -104.7517
9.8862 | -104.7451
9.8898 | -104.737
9.8934 | -104.7358
9.897 | -104.7325
9.9006 | -104.7301
9.9042 | -104.7412
9.9079 | -104.7441
9.9115 | -104.7524
9.9151 | -104.7614
9.9187 | -104.7716
9.9223 | -104.7827
9.9259 | -104.7949
9.9295 | -104.8018
9.9332 | -104.8077
9.9368 | -104.818
9.9404 | -104.8223
9.944 | -104.8193
9.9476 | -104.8156
9.9512 | -104.8137
9.9548 | -104.8077
9.9584 | -104.8009
9.9621 | -104.7948
9.9657 | -104.7859
9.9693 | -104.7746
9.9729 | -104.7645
9.9765 | -104.7531
9.9801 | -104.7415
9.9837 | -104.7303
9.9874 | -104.7254
9.991 | -104.7263
9.9946 | -104.7297
9.9982 | -104.7383
10.0018 | -104.7513
10.0054 | -104.7654
10.009 | -104.7839
10.0126 | -104.8013
10.0163 | -104.8167
10.0199 | -104.8355
10.0235 | -104.8626
10.0271 | -104.889
10.0307 | -104.912
10.0343 | -104.9314
10.0379 | -104.9484
10.0416 | -104.9656
10.0452 | -104.9809
10.0488 | -104.993
10.0524 | -105.005
10.056 | -105.0102
10.0596 | -105.0116
10.0632 | -105.0115
10.0668 | -105.0117
10.0705 | -105.0137
10.0741 | -105.0115
10.0777 | -105.0098
10.0813 | -105.0066
10.0849 | -105.001
10.0885 | -105.0015
10.0921 | -105.0048
10.0958 | -105.0021
10.0994 | -104.9997
10.103 | -105.0019
10.1066 | -105.0044
10.1102 | -105.0048
10.1138 | -105.0055
10.1174 | -105.0105
10.121 | -105.0133
10.1247 | -105.0168
10.1283 | -105.0214
10.1319 | -105.0253
10.1355 | -105.0303
10.1391 | -105.0371
10.1427 | -105.038
10.1463 | -105.0408
10.15 | -105.0456
10.1536 | -105.0512
10.1572 | -105.0516
10.1608 | -105.0501
10.1644 | -105.0515
10.168 | -105.0486
10.1716 | -105.0412
10.1752 | -105.0365
10.1789 | -105.0235
10.1825 | -105.0124
10.1861 | -104.9965
10.1897 | -104.981
10.1933 | -104.9621
10.1969 | -104.9408
10.2005 | -104.9192
10.2042 | -104.8984
10.2078 | -104.874
10.2114 | -104.8506
10.215 | -104.831
10.2186 | -104.8116
10.2222 | -104.7942
10.2258 | -104.7802
10.2294 | -104.7713
10.2331 | -104.7643
10.2367 | -104.7584
10.2403 | -104.7618
10.2439 | -104.774
10.2475 | -104.7898
10.2511 | -104.809
10.2547 | -104.836
10.2584 | -104.8713
10.262 | -104.9118
10.2656 | -104.9559
10.2692 | -105.0043
10.2728 | -105.0549
10.2764 | -105.1078
10.28 | -105.1588
10.2836 | -105.2063
10.2873 | -105.2513
10.2909 | -105.2994
10.2945 | -105.3395
10.2981 | -105.3731
10.3017 | -105.4031
10.3053 | -105.4325
10.3089 | -105.4525
10.3126 | -105.4663
10.3162 | -105.4787
10.3198 | -105.4819
10.3234 | -105.4752
10.327 | -105.4729
10.3306 | -105.4621
10.3342 | -105.4428
10.3378 | -105.4143
10.3415 | -105.3855
10.3451 | -105.3547
10.3487 | -105.3172
10.3523 | -105.2789
10.3559 | -105.2442
10.3595 | -105.2096
10.3631 | -105.1745
10.3668 | -105.1418
10.3704 | -105.113
10.374 | -105.0869
10.3776 | -105.0636
10.3812 | -105.0456
10.3848 | -105.0345
10.3884 | -105.022
10.3921 | -105.0134
10.3957 | -105.0152
10.3993 | -105.0228
10.4029 | -105.0279
10.4065 | -105.0317
10.4101 | -105.0417
10.4137 | -105.0537
10.4173 | -105.0595
10.421 | -105.0669
10.4246 | -105.0715
10.4282 | -105.0748
10.4318 | -105.0796
10.4354 | -105.0821
10.439 | -105.0827
10.4426 | -105.0795
10.4463 | -105.077
10.4499 | -105.0697
10.4535 | -105.0624
10.4571 | -105.055
10.4607 | -105.0487
10.4643 | -105.0355
10.4679 | -105.0208
10.4715 | -105.0125
10.4752 | -105.0059
10.4788 | -104.9954
10.4824 | -104.987
10.486 | -104.9841
10.4896 | -104.9807
10.4932 | -104.9719
10.4968 | -104.9639
10.5005 | -104.9594
10.5041 | -104.9548
10.5077 | -104.949
10.5113 | -104.9425
10.5149 | -104.9322
10.5185 | -104.9263
10.5221 | -104.9248
10.5257 | -104.9257
10.5294 | -104.9251
10.533 | -104.9251
10.5366 | -104.9258
10.5402 | -104.9282
10.5438 | -104.9286
10.5474 | -104.9274
10.551 | -104.9261
10.5547 | -104.9293
10.5583 | -104.9356
10.5619 | -104.9354
10.5655 | -104.9326
10.5691 | -104.9358
10.5727 | -104.9337
10.5763 | -104.9289
10.5799 | -104.9205
10.5836 | -104.9142
10.5872 | -104.9093
10.5908 | -104.904
10.5944 | -104.8978
10.598 | -104.8908
10.6016 | -104.8848
10.6052 | -104.8809
10.6089 | -104.8776
10.6125 | -104.8763
10.6161 | -104.8793
10.6197 | -104.8837
10.6233 | -104.8877
10.6269 | -104.8915
10.6305 | -104.8983
10.6341 | -104.9116
10.6378 | -104.9257
10.6414 | -104.9401
10.645 | -104.958
10.6486 | -104.9772
10.6522 | -104.9972
10.6558 | -105.0172
10.6594 | -105.0343
10.6631 | -105.0532
10.6667 | -105.0724
10.6703 | -105.0939
10.6739 | -105.1098
10.6775 | -105.1176
10.6811 | -105.122
10.6847 | -105.1267
10.6883 | -105.1264
10.692 | -105.1262
10.6956 | -105.1276
10.6992 | -105.1275
10.7028 | -105.1208
10.7064 | -105.1095
10.71 | -105.0986
10.7136 | -105.0883
10.7173 | -105.0773
10.7209 | -105.0635
10.7245 | -105.0471
10.7281 | -105.0328
10.7317 | -105.0178
10.7353 | -105.0052
10.7389 | -104.9934
10.7425 | -104.9799
10.7462 | -104.9644
10.7498 | -104.9538
10.7534 | -104.9467
10.757 | -104.9387
10.7606 | -104.9301
10.7642 | -104.9189
10.7678 | -104.906
10.7715 | -104.8963
10.7751 | -104.8915
10.7787 | -104.8893
10.7823 | -104.8869
10.7859 | -104.8864
10.7895 | -104.885
10.7931 | -104.8813
10.7967 | -104.8805
10.8004 | -104.8836
10.804 | -104.8921
10.8076 | -104.9019
10.8112 | -104.9113
10.8148 | -104.9262
10.8184 | -104.9427
10.822 | -104.963
10.8257 | -104.9876
10.8293 | -105.0159
10.8329 | -105.0498
10.8365 | -105.0857
10.8401 | -105.1214
10.8437 | -105.1588
10.8473 | -105.1924
10.8509 | -105.2244
10.8546 | -105.2541
10.8582 | -105.286
10.8618 | -105.3213
10.8654 | -105.3517
10.869 | -105.3787
10.8726 | -105.4046
10.8762 | -105.4278
10.8799 | -105.4475
10.8835 | -105.4642
10.8871 | -105.4746
10.8907 | -105.48
10.8943 | -105.4836
10.8979 | -105.4839
10.9015 | -105.4838
10.9051 | -105.4822
10.9088 | -105.4809
10.9124 | -105.4786
10.916 | -105.4759
10.9196 | -105.4745
10.9232 | -105.475
10.9268 | -105.4742
10.9304 | -105.4728
10.9341 | -105.4694
10.9377 | -105.4713
10.9413 | -105.4722
10.9449 | -105.4834
10.9485 | -105.4866
10.9521 | -105.4963
10.9557 | -105.5065
10.9594 | -105.5138
10.963 | -105.5189
10.9666 | -105.5231
10.9702 | -105.5302
10.9738 | -105.5322
10.9774 | -105.5311
10.981 | -105.5322
10.9846 | -105.5311
10.9883 | -105.5273
10.9919 | -105.5198
10.9955 | -105.5148
10.9991 | -105.5118
11.0027 | -105.5088
11.0063 | -105.5031
11.0099 | -105.4953
11.0136 | -105.4905
11.0172 | -105.4792
11.0208 | -105.4643
11.0244 | -105.4503
11.028 | -105.4377
11.0316 | -105.4259
11.0352 | -105.4142
11.0388 | -105.3999
11.0425 | -105.3842
11.0461 | -105.372
11.0497 | -105.361
11.0533 | -105.3488
11.0569 | -105.337
11.0605 | -105.3268
11.0641 | -105.3167
11.0678 | -105.3068
11.0714 | -105.3018
11.075 | -105.3001
11.0786 | -105.2945
11.0822 | -105.2892
11.0858 | -105.2852
11.0894 | -105.2844
11.093 | -105.2854
11.0967 | -105.2828
11.1003 | -105.2829
11.1039 | -105.2834
11.1075 | -105.2874
11.1111 | -105.2937
11.1147 | -105.3001
11.1183 | -105.3064
11.122 | -105.3146
11.1256 | -105.3245
11.1292 | -105.3312
11.1328 | -105.3348
11.1364 | -105.3384
11.14 | -105.3411
11.1436 | -105.3404
11.1472 | -105.3398
11.1509 | -105.3388
11.1545 | -105.3381
11.1581 | -105.3284
11.1617 | -105.3176
11.1653 | -105.3064
11.1689 | -105.2995
11.1725 | -105.2918
11.1762 | -105.2846
11.1798 | -105.2821
11.1834 | -105.284
11.187 | -105.286
11.1906 | -105.2912
11.1942 | -105.3004
11.1978 | -105.3124
11.2014 | -105.3286
11.2051 | -105.3452
11.2087 | -105.3617
11.2123 | -105.381
11.2159 | -105.404
11.2195 | -105.4287
11.2231 | -105.4512
11.2267 | -105.4701
11.2304 | -105.485
11.234 | -105.5007
11.2376 | -105.5184
11.2412 | -105.5362
11.2448 | -105.5498
11.2484 | -105.5554
11.252 | -105.5579
11.2556 | -105.5593
11.2593 | -105.5593
11.2629 | -105.556
11.2665 | -105.5463
11.2701 | -105.5311
11.2737 | -105.5181
11.2773 | -105.5023
11.2809 | -105.4864
11.2846 | -105.4695
11.2882 | -105.4501
11.2918 | -105.4377
11.2954 | -105.425
11.299 | -105.4112
11.3026 | -105.3976
11.3062 | -105.3848
11.3098 | -105.374
11.3135 | -105.3675
11.3171 | -105.3638
11.3207 | -105.3652
11.3243 | -105.3656
11.3279 | -105.3678
11.3315 | -105.3667
11.3351 | -105.3711
11.3388 | -105.3738
11.3424 | -105.3765
11.346 | -105.3781
11.3496 | -105.379
11.3532 | -105.3772
11.3568 | -105.3773
11.3604 | -105.3766
11.364 | -105.3763
11.3677 | -105.3781
11.3713 | -105.3778
11.3749 | -105.374
11.3785 | -105.3618
11.3821 | -105.3457
11.3857 | -105.3327
11.3893 | -105.3231
11.393 | -105.3152
11.3966 | -105.3002
11.4002 | -105.2828
11.4038 | -105.2682
11.4074 | -105.2537
11.411 | -105.242
11.4146 | -105.2248
11.4182 | -105.2064
11.4219 | -105.19
11.4255 | -105.1762
11.4291 | -105.1636
11.4327 | -105.1524
11.4363 | -105.1414
11.4399 | -105.1328
11.4435 | -105.1244
11.4472 | -105.1174
11.4508 | -105.1145
11.4544 | -105.1086
11.458 | -105.1027
11.4616 | -105.0984
11.4652 | -105.088
11.4688 | -105.0788
11.4724 | -105.0709
11.4761 | -105.0645
11.4797 | -105.0569
11.4833 | -105.0475
11.4869 | -105.0384
11.4905 | -105.0309
11.4941 | -105.0248
11.4977 | -105.022
11.5014 | -105.0153
11.505 | -105.013
11.5086 | -105.0112
11.5122 | -105.0093
11.5158 | -105.0142
11.5194 | -105.0199
11.523 | -105.0243
11.5266 | -105.0316
11.5303 | -105.0398
11.5339 | -105.0502
11.5375 | -105.0648
11.5411 | -105.0828
11.5447 | -105.1057
11.5483 | -105.1309
11.5519 | -105.1542
11.5556 | -105.1761
11.5592 | -105.1997
11.5628 | -105.2237
11.5664 | -105.2482
11.57 | -105.2724
11.5736 | -105.2986
11.5772 | -105.3252
11.5808 | -105.3563
11.5845 | -105.3848
11.5881 | -105.4112
11.5917 | -105.4312
11.5953 | -105.4536
11.5989 | -105.4772
11.6025 | -105.4973
11.6061 | -105.5155
11.6098 | -105.5312
11.6134 | -105.5452
11.617 | -105.554
11.6206 | -105.5601
11.6242 | -105.5668
11.6278 | -105.5697
11.6314 | -105.5736
11.635 | -105.5787
11.6387 | -105.5815
11.6423 | -105.5837
11.6459 | -105.5853
11.6495 | -105.5881
11.6531 | -105.5924
11.6567 | -105.5987
11.6603 | -105.6058
11.664 | -105.6114
11.6676 | -105.6184
11.6712 | -105.6306
11.6748 | -105.6435
11.6784 | -105.6593
11.682 | -105.6746
11.6856 | -105.6876
11.6892 | -105.7037
11.6929 | -105.7216
11.6965 | -105.7413
11.7001 | -105.7592
11.7037 | -105.7761
11.7073 | -105.7902
11.7109 | -105.7976
11.7145 | -105.8062
11.7182 | -105.8181
11.7218 | -105.8247
11.7254 | -105.8307
11.729 | -105.831
11.7326 | -105.8303
11.7362 | -105.8275
11.7398 | -105.8189
11.7435 | -105.8056
11.7471 | -105.791
11.7507 | -105.7722
11.7543 | -105.753
11.7579 | -105.727
11.7615 | -105.7042
11.7651 | -105.6787
11.7687 | -105.6502
11.7724 | -105.6236
11.776 | -105.5973
11.7796 | -105.5739
11.7832 | -105.5503
11.7868 | -105.5322
11.7904 | -105.5179
11.794 | -105.5016
11.7977 | -105.4904
11.8013 | -105.4788
11.8049 | -105.4666
11.8085 | -105.4553
11.8121 | -105.4459
11.8157 | -105.4392
11.8193 | -105.4334
11.8229 | -105.4294
11.8266 | -105.4256
11.8302 | -105.4225
11.8338 | -105.4211
11.8374 | -105.4209
11.841 | -105.4174
11.8446 | -105.4141
11.8482 | -105.4128
11.8519 | -105.4106
11.8555 | -105.4086
11.8591 | -105.4072
11.8627 | -105.4054
11.8663 | -105.4038
11.8699 | -105.4025
11.8735 | -105.3991
11.8771 | -105.4016
11.8808 | -105.4083
11.8844 | -105.4119
11.888 | -105.4134
11.8916 | -105.4191
11.8952 | -105.4215
11.8988 | -105.4289
11.9024 | -105.4383
11.9061 | -105.4458
11.9097 | -105.4588
11.9133 | -105.4723
11.9169 | -105.4862
11.9205 | -105.4984
11.9241 | -105.5085
11.9277 | -105.5179
11.9313 | -105.5259
11.935 | -105.5353
11.9386 | -105.5445
11.9422 | -105.5517
11.9458 | -105.5589
11.9494 | -105.5655
11.953 | -105.573
11.9566 | -105.5766
11.9603 | -105.5805
11.9639 | -105.5864
11.9675 | -105.5903
11.9711 | -105.5952
11.9747 | -105.601
11.9783 | -105.6081
11.9819 | -105.6118
11.9855 | -105.6141
11.9892 | -105.6149
11.9928 | -105.6173
11.9964 | -105.6212
12 | -105.622
12.0036 | -105.6255
12.0072 | -105.6285
12.0108 | -105.6287
12.0145 | -105.6283
12.0181 | -105.6279
12.0217 | -105.625
12.0253 | -105.6216
12.0289 | -105.6188
12.0325 | -105.6162
12.0361 | -105.6146
12.0397 | -105.6101
12.0434 | -105.6052
12.047 | -105.6013
12.0506 | -105.5928
12.0542 | -105.5826
12.0578 | -105.573
12.0614 | -105.5638
12.065 | -105.5555
12.0687 | -105.5506
12.0723 | -105.5446
12.0759 | -105.538
12.0795 | -105.5336
12.0831 | -105.5307
12.0867 | -105.5288
12.0903 | -105.5281
12.0939 | -105.5345
12.0976 | -105.5419
12.1012 | -105.5541
12.1048 | -105.5654
12.1084 | -105.5772
12.112 | -105.593
12.1156 | -105.6115
12.1192 | -105.6308
12.1229 | -105.6528
12.1265 | -105.6762
12.1301 | -105.6989
12.1337 | -105.7247
12.1373 | -105.7516
12.1409 | -105.7782
12.1445 | -105.8055
12.1481 | -105.831
12.1518 | -105.8557
12.1554 | -105.8804
12.159 | -105.9045
12.1626 | -105.9256
12.1662 | -105.9427
12.1698 | -105.9571
12.1734 | -105.9674
12.1771 | -105.9808
12.1807 | -105.9856
12.1843 | -105.9858
12.1879 | -105.9857
12.1915 | -105.9866
12.1951 | -105.9867
12.1987 | -105.9849
12.2023 | -105.9827
12.206 | -105.9755
12.2096 | -105.9676
12.2132 | -105.9575
12.2168 | -105.9502
12.2204 | -105.9413
12.224 | -105.9333
12.2276 | -105.9244
12.2313 | -105.9132
12.2349 | -105.9024
12.2385 | -105.8902
12.2421 | -105.8785
12.2457 | -105.8656
12.2493 | -105.8545
12.2529 | -105.8482
12.2565 | -105.8422
12.2602 | -105.8338
12.2638 | -105.8284
12.2674 | -105.825
12.271 | -105.819
12.2746 | -105.8144
12.2782 | -105.8111
12.2818 | -105.8091
12.2855 | -105.8024
12.2891 | -105.7934
12.2927 | -105.7876
12.2963 | -105.7811
12.2999 | -105.7727
12.3035 | -105.7642
12.3071 | -105.7567
12.3108 | -105.7458
12.3144 | -105.7347
12.318 | -105.7247
12.3216 | -105.711
12.3252 | -105.6976
12.3288 | -105.6871
12.3324 | -105.6745
12.336 | -105.6606
12.3397 | -105.6501
12.3433 | -105.6406
12.3469 | -105.63
12.3505 | -105.6204
12.3541 | -105.6097
12.3577 | -105.6033
12.3613 | -105.6
12.365 | -105.5937
12.3686 | -105.5883
12.3722 | -105.5839
12.3758 | -105.58
12.3794 | -105.5766
12.383 | -105.5726
12.3866 | -105.5655
12.3902 | -105.5594
12.3939 | -105.5516
12.3975 | -105.5463
12.4011 | -105.5409
12.4047 | -105.537
12.4083 | -105.5375
12.4119 | -105.5344
12.4155 | -105.5339
12.4192 | -105.5325
12.4228 | -105.5298
12.4264 | -105.5263
12.43 | -105.5204
12.4336 | -105.5185
12.4372 | -105.5218
12.4408 | -105.5276
12.4444 | -105.5346
12.4481 | -105.5423
12.4517 | -105.5528
12.4553 | -105.5619
12.4589 | -105.5732
12.4625 | -105.5886
12.4661 | -105.6028
12.4697 | -105.6137
12.4734 | -105.6261
12.477 | -105.6374
12.4806 | -105.6487
12.4842 | -105.6608
12.4878 | -105.6716
12.4914 | -105.6835
12.495 | -105.6948
12.4986 | -105.7082
12.5023 | -105.7215
12.5059 | -105.7341
12.5095 | -105.75
12.5131 | -105.7603
12.5167 | -105.7722
12.5203 | -105.7828
12.5239 | -105.7955
12.5276 | -105.8075
12.5312 | -105.8155
12.5348 | -105.8188
12.5384 | -105.8218
12.542 | -105.826
12.5456 | -105.8293
12.5492 | -105.8294
12.5528 | -105.8247
12.5565 | -105.8181
12.5601 | -105.8094
12.5637 | -105.8004
12.5673 | -105.7881
12.5709 | -105.7709
12.5745 | -105.7546
12.5781 | -105.7369
12.5818 | -105.7154
12.5854 | -105.6935
12.589 | -105.669
12.5926 | -105.644
12.5962 | -105.6209
12.5998 | -105.5982
12.6034 | -105.5798
12.607 | -105.5639
12.6107 | -105.5476
12.6143 | -105.5306
12.6179 | -105.5145
12.6215 | -105.4983
12.6251 | -105.4865
12.6287 | -105.4812
12.6323 | -105.4795
12.636 | -105.4799
12.6396 | -105.4814
12.6432 | -105.4844
12.6468 | -105.4912
12.6504 | -105.5008
12.654 | -105.5123
12.6576 | -105.5243
12.6612 | -105.5343
12.6649 | -105.5465
12.6685 | -105.5605
12.6721 | -105.5734
12.6757 | -105.585
12.6793 | -105.5978
12.6829 | -105.6094
12.6865 | -105.6166
12.6902 | -105.6216
12.6938 | -105.63
12.6974 | -105.6413
12.701 | -105.6504
12.7046 | -105.6596
12.7082 | -105.6689
12.7118 | -105.6773
12.7154 | -105.6855
12.7191 | -105.6936
12.7227 | -105.7005
12.7263 | -105.7035
12.7299 | -105.7068
12.7335 | -105.7115
12.7371 | -105.7174
12.7407 | -105.7258
12.7444 | -105.735
12.748 | -105.7417
12.7516 | -105.7503
12.7552 | -105.7604
12.7588 | -105.7675
12.7624 | -105.7713
12.766 | -105.7783
12.7696 | -105.7852
12.7733 | -105.7937
12.7769 | -105.8041
12.7805 | -105.8103
12.7841 | -105.8185
12.7877 | -105.8245
12.7913 | -105.8293
12.7949 | -105.8375
12.7986 | -105.8442
12.8022 | -105.8484
12.8058 | -105.8534
12.8094 | -105.8569
12.813 | -105.8571
12.8166 | -105.855
12.8202 | -105.855
12.8238 | -105.8557
12.8275 | -105.8542
12.8311 | -105.8531
12.8347 | -105.8493
12.8383 | -105.8451
12.8419 | -105.8423
12.8455 | -105.8392
12.8491 | -105.8325
12.8528 | -105.8239
12.8564 | -105.8172
12.86 | -105.8112
12.8636 | -105.8062
12.8672 | -105.8029
12.8708 | -105.8047
12.8744 | -105.8082
12.878 | -105.8092
12.8817 | -105.8125
12.8853 | -105.8184
12.8889 | -105.8251
12.8925 | -105.8326
12.8961 | -105.8424
12.8997 | -105.858
12.9033 | -105.8769
12.907 | -105.8936
12.9106 | -105.9108
12.9142 | -105.9267
12.9178 | -105.9406
12.9214 | -105.9554
12.925 | -105.9684
12.9286 | -105.9815
12.9322 | -105.9945
12.9359 | -106.0084
12.9395 | -106.0227
12.9431 | -106.0392
12.9467 | -106.0525
12.9503 | -106.0597
12.9539 | -106.0678
12.9575 | -106.0758
12.9612 | -106.0824
12.9648 | -106.0878
12.9684 | -106.0918
12.972 | -106.0949
12.9756 | -106.0964
12.9792 | -106.0908
12.9828 | -106.0832
12.9864 | -106.0747
12.9901 | -106.0652
12.9937 | -106.0516
12.9973 | -106.0377
13.0009 | -106.0198
13.0045 | -105.9996
13.0081 | -105.9774
13.0117 | -105.9551
13.0154 | -105.9297
13.019 | -105.9063
13.0226 | -105.8821
13.0262 | -105.8553
13.0298 | -105.8317
13.0334 | -105.808
13.037 | -105.7868
13.0406 | -105.7703
13.0443 | -105.7547
13.0479 | -105.7403
13.0515 | -105.7239
13.0551 | -105.7095
13.0587 | -105.6937
13.0623 | -105.6771
13.0659 | -105.6661
13.0696 | -105.6559
13.0732 | -105.6465
13.0768 | -105.6399
13.0804 | -105.6354
13.084 | -105.6352
13.0876 | -105.6351
13.0912 | -105.6361
13.0949 | -105.6383
13.0985 | -105.6414
13.1021 | -105.6483
13.1057 | -105.6533
13.1093 | -105.6579
13.1129 | -105.6619
13.1165 | -105.6665
13.1201 | -105.6749
13.1238 | -105.6811
13.1274 | -105.685
13.131 | -105.6887
13.1346 | -105.6939
13.1382 | -105.7015
13.1418 | -105.7096
13.1454 | -105.7182
13.1491 | -105.7275
13.1527 | -105.7384
13.1563 | -105.747
13.1599 | -105.757
13.1635 | -105.7688
13.1671 | -105.7847
13.1707 | -105.8005
13.1743 | -105.8148
13.178 | -105.8279
13.1816 | -105.8402
13.1852 | -105.8517
13.1888 | -105.865
13.1924 | -105.878
13.196 | -105.8883
13.1996 | -105.8984
13.2033 | -105.9075
13.2069 | -105.9115
13.2105 | -105.9171
13.2141 | -105.9199
13.2177 | -105.9198
13.2213 | -105.915
13.2249 | -105.9081
13.2285 | -105.8983
13.2322 | -105.8823
13.2358 | -105.8609
13.2394 | -105.8384
13.243 | -105.8144
13.2466 | -105.791
13.2502 | -105.7693
13.2538 | -105.7468
13.2575 | -105.7214
13.2611 | -105.6988
13.2647 | -105.6734
13.2683 | -105.6442
13.2719 | -105.6148
13.2755 | -105.5845
13.2791 | -105.558
13.2827 | -105.5349
13.2864 | -105.5141
13.29 | -105.4976
13.2936 | -105.481
13.2972 | -105.4683
13.3008 | -105.4576
13.3044 | -105.4496
13.308 | -105.4429
13.3117 | -105.4371
13.3153 | -105.4337
13.3189 | -105.4312
13.3225 | -105.4307
13.3261 | -105.4325
13.3297 | -105.4351
13.3333 | -105.4396
13.3369 | -105.4448
13.3406 | -105.4491
13.3442 | -105.4574
13.3478 | -105.4699
13.3514 | -105.4833
13.355 | -105.4967
13.3586 | -105.5098
13.3622 | -105.5233
13.3659 | -105.5378
13.3695 | -105.5523
13.3731 | -105.57
13.3767 | -105.5882
13.3803 | -105.6078
13.3839 | -105.6263
13.3875 | -105.6417
13.3911 | -105.6563
13.3948 | -105.6702
13.3984 | -105.6835
13.402 | -105.694
13.4056 | -105.7087
13.4092 | -105.7219
13.4128 | -105.7317
13.4164 | -105.7423
13.4201 | -105.754
13.4237 | -105.768
13.4273 | -105.7835
13.4309 | -105.7959
13.4345 | -105.8081
13.4381 | -105.8184
13.4417 | -105.8249
13.4453 | -105.8343
13.449 | -105.8446
13.4526 | -105.8522
13.4562 | -105.863
13.4598 | -105.874
13.4634 | -105.8823
13.467 | -105.8907
13.4706 | -105.896
13.4743 | -105.8992
13.4779 | -105.9023
13.4815 | -105.9022
13.4851 | -105.901
13.4887 | -105.8988
13.4923 | -105.8953
13.4959 | -105.8899
13.4995 | -105.8839
13.5032 | -105.8797
13.5068 | -105.875
13.5104 | -105.8722
13.514 | -105.8669
13.5176 | -105.8627
13.5212 | -105.8569
13.5248 | -105.8459
13.5285 | -105.8347
13.5321 | -105.8255
13.5357 | -105.8134
13.5393 | -105.8015
13.5429 | -105.7917
13.5465 | -105.7832
13.5501 | -105.7746
13.5537 | -105.7679
13.5574 | -105.762
13.561 | -105.7567
13.5646 | -105.7521
13.5682 | -105.7472
13.5718 | -105.7429
13.5754 | -105.7396
13.579 | -105.7385
13.5827 | -105.7413
13.5863 | -105.7433
13.5899 | -105.7475
13.5935 | -105.753
13.5971 | -105.7565
13.6007 | -105.7599
13.6043 | -105.7642
13.6079 | -105.7715
13.6116 | -105.7799
13.6152 | -105.7851
13.6188 | -105.7888
13.6224 | -105.7939
13.626 | -105.8038
13.6296 | -105.8169
13.6332 | -105.829
13.6369 | -105.8414
13.6405 | -105.8537
13.6441 | -105.8636
13.6477 | -105.8741
13.6513 | -105.8856
13.6549 | -105.8976
13.6585 | -105.9116
13.6622 | -105.9219
13.6658 | -105.93
13.6694 | -105.9402
13.673 | -105.9509
13.6766 | -105.9613
13.6802 | -105.9704
13.6838 | -105.9793
13.6874 | -105.9895
13.6911 | -105.9986
13.6947 | -106.0055
13.6983 | -106.0109
13.7019 | -106.0161
13.7055 | -106.0233
13.7091 | -106.0314
13.7127 | -106.0391
13.7164 | -106.0497
13.72 | -106.0631
13.7236 | -106.0766
13.7272 | -106.0888
13.7308 | -106.1009
13.7344 | -106.114
13.738 | -106.1236
13.7416 | -106.1337
13.7453 | -106.1464
13.7489 | -106.161
13.7525 | -106.1699
13.7561 | -106.1754
13.7597 | -106.1821
13.7633 | -106.1884
13.7669 | -106.19
13.7706 | -106.1929
13.7742 | -106.1935
13.7778 | -106.1922
13.7814 | -106.1924
13.785 | -106.1932
13.7886 | -106.1906
13.7922 | -106.1884
13.7958 | -106.185
13.7995 | -106.1832
13.8031 | -106.179
13.8067 | -106.1745
13.8103 | -106.1684
13.8139 | -106.1651
13.8175 | -106.1647
13.8211 | -106.1643
13.8248 | -106.1632
13.8284 | -106.1608
13.832 | -106.1593
13.8356 | -106.1605
13.8392 | -106.1616
13.8428 | -106.1673
13.8464 | -106.1744
13.85 | -106.1816
13.8537 | -106.1902
13.8573 | -106.1967
13.8609 | -106.2037
13.8645 | -106.2141
13.8681 | -106.2199
13.8717 | -106.2223
13.8753 | -106.226
13.879 | -106.2327
13.8826 | -106.2375
13.8862 | -106.24
13.8898 | -106.2444
13.8934 | -106.2478
13.897 | -106.2496
13.9006 | -106.2534
13.9042 | -106.2549
13.9079 | -106.2553
13.9115 | -106.2553
13.9151 | -106.2493
13.9187 | -106.2436
13.9223 | -106.2423
13.9259 | -106.2422
13.9295 | -106.2386
13.9332 | -106.2335
13.9368 | -106.2289
13.9404 | -106.226
13.944 | -106.2185
13.9476 | -106.211
13.9512 | -106.203
13.9548 | -106.1946
13.9584 | -106.1848
13.9621 | -106.1706
13.9657 | -106.1553
13.9693 | -106.1392
13.9729 | -106.1192
13.9765 | -106.1037
13.9801 | -106.0885
13.9837 | -106.0714
13.9874 | -106.0561
13.991 | -106.0418
13.9946 | -106.0279
13.9982 | -106.0138
14.0018 | -106.001
14.0054 | -105.9889
14.009 | -105.9779
14.0126 | -105.966
14.0163 | -105.9569
14.0199 | -105.9484
14.0235 | -105.9401
14.0271 | -105.9313
14.0307 | -105.9266
14.0343 | -105.9253
14.0379 | -105.9252
14.0416 | -105.9219
14.0452 | -105.9236
14.0488 | -105.927
14.0524 | -105.9344
14.056 | -105.9456
14.0596 | -105.9595
14.0632 | -105.9713
14.0668 | -105.9874
14.0705 | -106.0092
14.0741 | -106.0371
14.0777 | -106.0605
14.0813 | -106.0838
14.0849 | -106.1099
14.0885 | -106.1372
14.0921 | -106.1666
14.0958 | -106.1974
14.0994 | -106.2283
14.103 | -106.2598
14.1066 | -106.2881
14.1102 | -106.3137
14.1138 | -106.3404
14.1174 | -106.3652
14.121 | -106.3839
14.1247 | -106.3982
14.1283 | -106.415
14.1319 | -106.4296
14.1355 | -106.4406
14.1391 | -106.4489
14.1427 | -106.4511
14.1463 | -106.4524
14.15 | -106.4543
14.1536 | -106.453
14.1572 | -106.4512
14.1608 | -106.4473
14.1644 | -106.4446
14.168 | -106.4425
14.1716 | -106.4369
14.1752 | -106.4329
14.1789 | -106.4217
14.1825 | -106.4112
14.1861 | -106.4051
14.1897 | -106.4002
14.1933 | -106.3945
14.1969 | -106.3924
14.2005 | -106.3893
14.2042 | -106.39
14.2078 | -106.3917
14.2114 | -106.3906
14.215 | -106.3909
14.2186 | -106.3901
14.2222 | -106.3937
14.2258 | -106.3984
14.2294 | -106.4062
14.2331 | -106.4118
14.2367 | -106.4171
14.2403 | -106.4223
14.2439 | -106.4299
14.2475 | -106.4368
14.2511 | -106.4431
14.2547 | -106.4494
14.2584 | -106.457
14.262 | -106.4616
14.2656 | -106.4703
14.2692 | -106.4782
14.2728 | -106.4841
14.2764 | -106.4917
14.28 | -106.4982
14.2836 | -106.5022
14.2873 | -106.5084
14.2909 | -106.5124
14.2945 | -106.5161
14.2981 | -106.5178
14.3017 | -106.5187
14.3053 | -106.5203
14.3089 | -106.5202
14.3126 | -106.5222
14.3162 | -106.525
14.3198 | -106.526
14.3234 | -106.5287
14.327 | -106.5296
14.3306 | -106.5242
14.3342 | -106.5131
14.3378 | -106.5045
14.3415 | -106.4981
14.3451 | -106.4902
14.3487 | -106.4825
14.3523 | -106.4725
14.3559 | -106.46
14.3595 | -106.4482
14.3631 | -106.4358
14.3668 | -106.4225
14.3704 | -106.4094
14.374 | -106.397
14.3776 | -106.3841
14.3812 | -106.3722
14.3848 | -106.3628
14.3884 | -106.3514
14.3921 | -106.3384
14.3957 | -106.3262
14.3993 | -106.3139
14.4029 | -106.3028
14.4065 | -106.291
14.4101 | -106.2804
14.4137 | -106.274
14.4173 | -106.2691
14.421 | -106.2647
14.4246 | -106.2574
14.4282 | -106.2498
14.4318 | -106.2431
14.4354 | -106.2358
14.439 | -106.2283
14.4426 | -106.222
14.4463 | -106.2181
14.4499 | -106.2135
14.4535 | -106.2067
14.4571 | -106.2002
14.4607 | -106.1914
14.4643 | -106.1835
14.4679 | -106.179
14.4715 | -106.1728
14.4752 | -106.1668
14.4788 | -106.16
14.4824 | -106.1548
14.486 | -106.1528
14.4896 | -106.1482
14.4932 | -106.1413
14.4968 | -106.1383
14.5005 | -106.1362
14.5041 | -106.1375
14.5077 | -106.1394
14.5113 | -106.1427
14.5149 | -106.1469
14.5185 | -106.1537
14.5221 | -106.1635
14.5257 | -106.1724
14.5294 | -106.184
14.533 | -106.1976
14.5366 | -106.2144
14.5402 | -106.2317
14.5438 | -106.2475
14.5474 | -106.2623
14.551 | -106.2764
14.5547 | -106.2928
14.5583 | -106.3132
14.5619 | -106.3336
14.5655 | -106.3513
14.5691 | -106.3677
14.5727 | -106.3813
14.5763 | -106.3942
14.5799 | -106.4061
14.5836 | -106.4197
14.5872 | -106.4305
14.5908 | -106.4379
14.5944 | -106.4315
14.598 | -106.4324
14.6016 | -106.434
14.6052 | -106.4328
14.6089 | -106.432
14.6125 | -106.4271
14.6161 | -106.4216
14.6197 | -106.413
14.6233 | -106.3995
14.6269 | -106.3858
14.6305 | -106.3711
14.6341 | -106.3571
14.6378 | -106.3427
14.6414 | -106.3296
14.645 | -106.3155
14.6486 | -106.3002
14.6522 | -106.2855
14.6558 | -106.2722
14.6594 | -106.2596
14.6631 | -106.2486
14.6667 | -106.2398
14.6703 | -106.2339
14.6739 | -106.2265
14.6775 | -106.2204
14.6811 | -106.2153
14.6847 | -106.2115
14.6883 | -106.2102
14.692 | -106.2076
14.6956 | -106.2042
14.6992 | -106.2029
14.7028 | -106.2058
14.7064 | -106.2095
14.71 | -106.2148
14.7136 | -106.2185
14.7173 | -106.2229
14.7209 | -106.2265
14.7245 | -106.2305
14.7281 | -106.2357
14.7317 | -106.2409
14.7353 | -106.2484
14.7389 | -106.2535
14.7425 | -106.2589
14.7462 | -106.2635
14.7498 | -106.2703
14.7534 | -106.2782
14.757 | -106.2887
14.7606 | -106.3005
14.7642 | -106.3101
14.7678 | -106.3171
14.7715 | -106.3223
14.7751 | -106.332
14.7787 | -106.3421
14.7823 | -106.3548
14.7859 | -106.3664
14.7895 | -106.3755
14.7931 | -106.3814
14.7967 | -106.3882
14.8004 | -106.3977
14.804 | -106.4094
14.8076 | -106.4188
14.8112 | -106.4303
14.8148 | -106.4398
14.8184 | -106.4477
14.822 | -106.4537
14.8257 | -106.4592
14.8293 | -106.4653
14.8329 | -106.4693
14.8365 | -106.4741
14.8401 | -106.4806
14.8437 | -106.4848
14.8473 | -106.4875
14.8509 | -106.491
14.8546 | -106.4947
14.8582 | -106.4984
14.8618 | -106.5025
14.8654 | -106.5048
14.869 | -106.5063
14.8726 | -106.5037
14.8762 | -106.4994
14.8799 | -106.4962
14.8835 | -106.4942
14.8871 | -106.494
14.8907 | -106.491
14.8943 | -106.4895
14.8979 | -106.4902
14.9015 | -106.4863
14.9051 | -106.4826
14.9088 | -106.4824
14.9124 | -106.4829
14.916 | -106.4791
14.9196 | -106.4732
14.9232 | -106.4665
14.9268 | -106.4619
14.9304 | -106.4592
14.9341 | -106.4564
14.9377 | -106.4528
14.9413 | -106.4496
14.9449 | -106.4436
14.9485 | -106.435
14.9521 | -106.4308
14.9557 | -106.4259
14.9594 | -106.4231
14.963 | -106.4188
14.9666 | -106.4117
14.9702 | -106.4043
14.9738 | -106.3979
14.9774 | -106.3884
14.981 | -106.3826
14.9846 | -106.3782
14.9883 | -106.3716
14.9919 | -106.3644
14.9955 | -106.3551
14.9991 | -106.3485
15.0027 | -106.3447
15.0063 | -106.3389
15.0099 | -106.3341
15.0136 | -106.3274
15.0172 | -106.3208
15.0208 | -106.3155
15.0244 | -106.3094
15.028 | -106.3028
15.0316 | -106.2969
15.0352 | -106.2919
15.0388 | -106.2874
15.0425 | -106.28
15.0461 | -106.2764
15.0497 | -106.2714
15.0533 | -106.2637
15.0569 | -106.2575
15.0605 | -106.2535
15.0641 | -106.2491
15.0678 | -106.2439
15.0714 | -106.2392
15.075 | -106.2376
15.0786 | -106.2345
15.0822 | -106.231
15.0858 | -106.2339
15.0894 | -106.2375
15.093 | -106.2395
15.0967 | -106.2389
15.1003 | -106.2394
15.1039 | -106.2433
15.1075 | -106.2467
15.1111 | -106.2509
15.1147 | -106.2612
15.1183 | -106.2703
15.122 | -106.2788
15.1256 | -106.2878
15.1292 | -106.2972
15.1328 | -106.3055
15.1364 | -106.3166
15.14 | -106.3293
15.1436 | -106.341
15.1472 | -106.351
15.1509 | -106.3607
15.1545 | -106.3743
15.1581 | -106.3859
15.1617 | -106.4014
15.1653 | -106.4179
15.1689 | -106.4314
15.1725 | -106.444
15.1762 | -106.4548
15.1798 | -106.4658
15.1834 | -106.4786
15.187 | -106.4883
15.1906 | -106.4954
15.1942 | -106.5052
15.1978 | -106.5142
15.2014 | -106.5213
15.2051 | -106.5275
15.2087 | -106.5355
15.2123 | -106.5408
15.2159 | -106.5436
15.2195 | -106.5424
15.2231 | -106.5434
15.2267 | -106.5419
15.2304 | -106.5373
15.234 | -106.5294
15.2376 | -106.5222
15.2412 | -106.5153
15.2448 | -106.5098
15.2484 | -106.502
15.252 | -106.4926
15.2556 | -106.4824
15.2593 | -106.4727
15.2629 | -106.4581
15.2665 | -106.443
15.2701 | -106.4321
15.2737 | -106.4185
15.2773 | -106.4061
15.2809 | -106.3941
15.2846 | -106.3814
15.2882 | -106.3708
15.2918 | -106.3606
15.2954 | -106.3551
15.299 | -106.3485
15.3026 | -106.3424
15.3062 | -106.3395
15.3098 | -106.3354
15.3135 | -106.3324
15.3171 | -106.3293
15.3207 | -106.3267
15.3243 | -106.3238
15.3279 | -106.3216
15.3315 | -106.3236
15.3351 | -106.3264
15.3388 | -106.329
15.3424 | -106.3345
15.346 | -106.338
15.3496 | -106.3434
15.3532 | -106.3523
15.3568 | -106.3612
15.3604 | -106.3709
15.364 | -106.3814
15.3677 | -106.3911
15.3713 | -106.3984
15.3749 | -106.4055
15.3785 | -106.4141
15.3821 | -106.423
15.3857 | -106.4327
15.3893 | -106.4421
15.393 | -106.4512
15.3966 | -106.461
15.4002 | -106.4689
15.4038 | -106.4761
15.4074 | -106.4847
15.411 | -106.494
15.4146 | -106.4991
15.4182 | -106.5015
15.4219 | -106.504
15.4255 | -106.5061
15.4291 | -106.5067
15.4327 | -106.5074
15.4363 | -106.5063
15.4399 | -106.5061
15.4435 | -106.5043
15.4472 | -106.5029
15.4508 | -106.505
15.4544 | -106.5074
15.458 | -106.509
15.4616 | -106.5107
15.4652 | -106.5106
15.4688 | -106.5134
15.4724 | -106.5182
15.4761 | -106.5244
15.4797 | -106.5301
15.4833 | -106.537
15.4869 | -106.5453
15.4905 | -106.5499
15.4941 | -106.5547
15.4977 | -106.5621
15.5014 | -106.5688
15.505 | -106.5798
15.5086 | -106.5878
15.5122 | -106.5925
15.5158 | -106.5956
15.5194 | -106.6021
15.523 | -106.6071
15.5266 | -106.6069
15.5303 | -106.608
15.5339 | -106.61
15.5375 | -106.6098
15.5411 | -106.6058
15.5447 | -106.6031
15.5483 | -106.5984
15.5519 | -106.5898
15.5556 | -106.5809
15.5592 | -106.571
15.5628 | -106.5596
15.5664 | -106.5482
15.57 | -106.5332
15.5736 | -106.5142
15.5772 | -106.4938
15.5808 | -106.4722
15.5845 | -106.4496
15.5881 | -106.4291
15.5917 | -106.4121
15.5953 | -106.3915
15.5989 | -106.3685
15.6025 | -106.3447
15.6061 | -106.3217
15.6098 | -106.3006
15.6134 | -106.2798
15.617 | -106.258
15.6206 | -106.2396
15.6242 | -106.2212
15.6278 | -106.2002
15.6314 | -106.182
15.635 | -106.1676
15.6387 | -106.1625
15.6423 | -106.1516
15.6459 | -106.1414
15.6495 | -106.1331
15.6531 | -106.1253
15.6567 | -106.1207
15.6603 | -106.1206
15.664 | -106.1212
15.6676 | -106.1219
15.6712 | -106.1239
15.6748 | -106.1255
15.6784 | -106.1296
15.682 | -106.1364
15.6856 | -106.1438
15.6892 | -106.1534
15.6929 | -106.1647
15.6965 | -106.1769
15.7001 | -106.1891
15.7037 | -106.2022
15.7073 | -106.2162
15.7109 | -106.2307
15.7145 | -106.2454
15.7182 | -106.2614
15.7218 | -106.2717
15.7254 | -106.279
15.729 | -106.2843
15.7326 | -106.2924
15.7362 | -106.3027
15.7398 | -106.3114
15.7435 | -106.3181
15.7471 | -106.3219
15.7507 | -106.3278
15.7543 | -106.333
15.7579 | -106.3332
15.7615 | -106.3328
15.7651 | -106.3345
15.7687 | -106.3326
15.7724 | -106.3291
15.776 | -106.325
15.7796 | -106.3196
15.7832 | -106.3123
15.7868 | -106.3048
15.7904 | -106.2975
15.794 | -106.2891
15.7977 | -106.2829
15.8013 | -106.2762
15.8049 | -106.2673
15.8085 | -106.2584
15.8121 | -106.2502
15.8157 | -106.2397
15.8193 | -106.2307
15.8229 | -106.2225
15.8266 | -106.2131
15.8302 | -106.2035
15.8338 | -106.1951
15.8374 | -106.1841
15.841 | -106.1754
15.8446 | -106.1676
15.8482 | -106.162
15.8519 | -106.1581
15.8555 | -106.1535
15.8591 | -106.1469
15.8627 | -106.1407
15.8663 | -106.1354
15.8699 | -106.1299
15.8735 | -106.1237
15.8771 | -106.1167
15.8808 | -106.1101
15.8844 | -106.1031
15.888 | -106.0939
15.8916 | -106.0877
15.8952 | -106.0842
15.8988 | -106.0809
15.9024 | -106.0779
15.9061 | -106.0737
15.9097 | -106.0673
15.9133 | -106.064
15.9169 | -106.0617
15.9205 | -106.0587
15.9241 | -106.0563
15.9277 | -106.0549
15.9313 | -106.0569
15.935 | -106.0562
15.9386 | -106.0561
15.9422 | -106.0598
15.9458 | -106.0656
15.9494 | -106.072
15.953 | -106.079
15.9566 | -106.0876
15.9603 | -106.0956
15.9639 | -106.1002
15.9675 | -106.1056
15.9711 | -106.1121
15.9747 | -106.1212
15.9783 | -106.1288
15.9819 | -106.1332
15.9855 | -106.1412
15.9892 | -106.1513
15.9928 | -106.1591
15.9964 | -106.1654
16 | -106.1729
16.0036 | -106.1807
16.0072 | -106.1875
16.0108 | -106.1965
16.0145 | -106.2026
16.0181 | -106.2079
16.0217 | -106.2162
16.0253 | -106.2208
16.0289 | -106.2266
16.0325 | -106.2331
16.0361 | -106.2378
16.0397 | -106.2401
16.0434 | -106.2429
16.047 | -106.2445
16.0506 | -106.2485
16.0542 | -106.2535
16.0578 | -106.2567
16.0614 | -106.259
16.065 | -106.2622
16.0687 | -106.2617
16.0723 | -106.2595
16.0759 | -106.2578
16.0795 | -106.256
16.0831 | -106.2539
16.0867 | -106.2535
16.0903 | -106.2512
16.0939 | -106.2484
16.0976 | -106.2442
16.1012 | -106.2399
16.1048 | -106.2345
16.1084 | -106.229
16.112 | -106.2238
16.1156 | -106.2188
16.1192 | -106.2129
16.1229 | -106.2053
16.1265 | -106.1987
16.1301 | -106.1937
16.1337 | -106.191
16.1373 | -106.1892
16.1409 | -106.1856
16.1445 | -106.1821
16.1481 | -106.1813
16.1518 | -106.1822
16.1554 | -106.1845
16.159 | -106.1915
16.1626 | -106.1955
16.1662 | -106.202
16.1698 | -106.2074
16.1734 | -106.2117
16.1771 | -106.2153
16.1807 | -106.2192
16.1843 | -106.2233
16.1879 | -106.2282
16.1915 | -106.2333
16.1951 | -106.2403
16.1987 | -106.2472
16.2023 | -106.2543
16.206 | -106.2609
16.2096 | -106.2686
16.2132 | -106.2763
16.2168 | -106.2807
16.2204 | -106.2829
16.224 | -106.2844
16.2276 | -106.2887
16.2313 | -106.2902
16.2349 | -106.2893
16.2385 | -106.2902
16.2421 | -106.2877
16.2457 | -106.2856
16.2493 | -106.283
16.2529 | -106.2781
16.2565 | -106.2748
16.2602 | -106.2672
16.2638 | -106.2599
16.2674 | -106.2524
16.271 | -106.2452
16.2746 | -106.2375
16.2782 | -106.2281
16.2818 | -106.2178
16.2855 | -106.2082
16.2891 | -106.1986
16.2927 | -106.188
16.2963 | -106.179
16.2999 | -106.172
16.3035 | -106.1635
16.3071 | -106.1553
16.3107 | -106.1476
16.3144 | -106.1424
16.318 | -106.1392
16.3216 | -106.1369
16.3252 | -106.1348
16.3288 | -106.1348
16.3324 | -106.1381
16.336 | -106.1433
16.3397 | -106.1493
16.3433 | -106.1553
16.3469 | -106.1623
16.3505 | -106.1701
16.3541 | -106.1794
16.3577 | -106.1902
16.3613 | -106.2036
16.365 | -106.2153
16.3686 | -106.2289
16.3722 | -106.2449
16.3758 | -106.2608
16.3794 | -106.2765
16.383 | -106.2909
16.3866 | -106.3066
16.3902 | -106.3251
16.3939 | -106.3424
16.3975 | -106.3578
16.4011 | -106.3774
16.4047 | -106.3975
16.4083 | -106.4154
16.4119 | -106.4335
16.4155 | -106.4527
16.4192 | -106.4731
16.4228 | -106.4905
16.4264 | -106.5073
16.43 | -106.524
16.4336 | -106.5395
16.4372 | -106.5534
16.4408 | -106.5652
16.4444 | -106.577
16.4481 | -106.5922
16.4517 | -106.6059
16.4553 | -106.6161
16.4589 | -106.6274
16.4625 | -106.6349
16.4661 | -106.6432
16.4697 | -106.6508
16.4734 | -106.6559
16.477 | -106.6638
16.4806 | -106.6718
16.4842 | -106.6775
16.4878 | -106.6838
16.4914 | -106.6875
16.495 | -106.6949
16.4986 | -106.7033
16.5023 | -106.7101
16.5059 | -106.7147
16.5095 | -106.7189
16.5131 | -106.7273
16.5167 | -106.735
16.5203 | -106.7422
16.5239 | -106.7499
16.5276 | -106.7604
16.5312 | -106.7705
16.5348 | -106.7788
16.5384 | -106.7882
16.542 | -106.7955
16.5456 | -106.8016
16.5492 | -106.8069
16.5528 | -106.8107
16.5565 | -106.8148
16.5601 | -106.8199
16.5637 | -106.8225
16.5673 | -106.8244
16.5709 | -106.8255
16.5745 | -106.824
16.5781 | -106.8229
16.5818 | -106.8226
16.5854 | -106.8191
16.589 | -106.8149
16.5926 | -106.8097
16.5962 | -106.8024
16.5998 | -106.8001
16.6034 | -106.7972
16.607 | -106.7918
16.6107 | -106.7877
16.6143 | -106.7822
16.6179 | -106.7791
16.6215 | -106.7772
16.6251 | -106.7713
16.6287 | -106.7683
16.6323 | -106.7668
16.636 | -106.7662
16.6396 | -106.7661
16.6432 | -106.7679
16.6468 | -106.7732
16.6504 | -106.7796
16.654 | -106.7862
16.6576 | -106.7925
16.6612 | -106.7988
16.6649 | -106.808
16.6685 | -106.8187
16.6721 | -106.8313
16.6757 | -106.8466
16.6793 | -106.867
16.6829 | -106.8834
16.6865 | -106.8975
16.6902 | -106.9129
16.6938 | -106.9306
16.6974 | -106.9494
16.701 | -106.9695
16.7046 | -106.9884
16.7082 | -107.0078
16.7118 | -107.0264
16.7154 | -107.0396
16.7191 | -107.0517
16.7227 | -107.0662
16.7263 | -107.0801
16.7299 | -107.0939
16.7335 | -107.1063
16.7371 | -107.1176
16.7407 | -107.1296
16.7444 | -107.141
16.748 | -107.1516
16.7516 | -107.1578
16.7552 | -107.1629
16.7588 | -107.166
16.7624 | -107.1668
16.766 | -107.1638
16.7696 | -107.1614
16.7733 | -107.1569
16.7769 | -107.1507
16.7805 | -107.1452
16.7841 | -107.1388
16.7877 | -107.1275
16.7913 | -107.1142
16.7949 | -107.1011
16.7986 | -107.0864
16.8022 | -107.0682
16.8058 | -107.0472
16.8094 | -107.0252
16.813 | -107.0051
16.8166 | -106.9837
16.8202 | -106.9619
16.8238 | -106.9419
16.8275 | -106.9207
16.8311 | -106.8987
16.8347 | -106.877
16.8383 | -106.8545
16.8419 | -106.8347
16.8455 | -106.8178
16.8491 | -106.8003
16.8528 | -106.7845
16.8564 | -106.7689
16.86 | -106.7531
16.8636 | -106.7418
16.8672 | -106.7316
16.8708 | -106.7207
16.8744 | -106.7083
16.878 | -106.7011
16.8817 | -106.6976
16.8853 | -106.6922
16.8889 | -106.6875
16.8925 | -106.6863
16.8961 | -106.6876
16.8997 | -106.6919
16.9033 | -106.6948
16.907 | -106.6982
16.9106 | -106.7008
16.9142 | -106.705
16.9178 | -106.7119
16.9214 | -106.7174
16.925 | -106.7245
16.9286 | -106.7335
16.9322 | -106.7436
16.9359 | -106.7557
16.9395 | -106.7671
16.9431 | -106.7782
16.9467 | -106.7884
16.9503 | -106.8002
16.9539 | -106.8119
16.9575 | -106.8254
16.9612 | -106.8403
16.9648 | -106.853
16.9684 | -106.8661
16.972 | -106.8783
16.9756 | -106.8881
16.9792 | -106.8981
16.9828 | -106.9084
16.9865 | -106.9178
16.9901 | -106.928
16.9937 | -106.9394
16.9973 | -106.9496
17.0009 | -106.9605
17.0045 | -106.9732
17.0081 | -106.986
17.0117 | -106.9991
17.0154 | -107.0131
17.019 | -107.0261
17.0226 | -107.0375
17.0262 | -107.0477
17.0298 | -107.0607
17.0334 | -107.0751
17.037 | -107.0895
17.0406 | -107.1036
17.0443 | -107.1189
17.0479 | -107.1347
17.0515 | -107.1485
17.0551 | -107.1595
17.0587 | -107.1708
17.0623 | -107.1811
17.0659 | -107.1895
17.0696 | -107.1996
17.0732 | -107.209
17.0768 | -107.2173
17.0804 | -107.2254
17.084 | -107.2301
17.0876 | -107.2326
17.0912 | -107.2365
17.0949 | -107.2406
17.0985 | -107.2436
17.1021 | -107.2467
17.1057 | -107.2469
17.1093 | -107.2411
17.1129 | -107.2382
17.1165 | -107.2376
17.1201 | -107.232
17.1238 | -107.2225
17.1274 | -107.2147
17.131 | -107.2061
17.1346 | -107.1941
17.1382 | -107.1812
17.1418 | -107.1677
17.1454 | -107.1546
17.1491 | -107.1425
17.1527 | -107.1297
17.1563 | -107.1164
17.1599 | -107.1035
17.1635 | -107.0891
17.1671 | -107.0726
17.1707 | -107.0562
17.1743 | -107.0411
17.178 | -107.0237
17.1816 | -107.0064
17.1852 | -106.9915
17.1888 | -106.9777
17.1924 | -106.9636
17.196 | -106.9496
17.1996 | -106.9375
17.2033 | -106.9243
17.2069 | -106.9099
17.2105 | -106.898
17.2141 | -106.8861
17.2177 | -106.8749
17.2213 | -106.8645
17.2249 | -106.8544
17.2285 | -106.8451
17.2322 | -106.8367
17.2358 | -106.8276
17.2394 | -106.8195
17.243 | -106.8144
17.2466 | -106.8102
17.2502 | -106.8073
17.2538 | -106.8043
17.2575 | -106.7985
17.2611 | -106.7924
17.2647 | -106.787
17.2683 | -106.7847
17.2719 | -106.787
17.2755 | -106.7865
17.2791 | -106.785
17.2827 | -106.7842
17.2864 | -106.783
17.29 | -106.785
17.2936 | -106.7859
17.2972 | -106.7856
17.3008 | -106.7853
17.3044 | -106.7839
17.308 | -106.7823
17.3117 | -106.7828
17.3153 | -106.7823
17.3189 | -106.7828
17.3225 | -106.7838
17.3261 | -106.7824
17.3297 | -106.783
17.3333 | -106.7831
17.3369 | -106.7815
17.3406 | -106.7791
17.3442 | -106.7798
17.3478 | -106.7801
17.3514 | -106.781
17.355 | -106.785
17.3586 | -106.7889
17.3622 | -106.7904
17.3659 | -106.7911
17.3695 | -106.7931
17.3731 | -106.7931
17.3767 | -106.7932
17.3803 | -106.7944
17.3839 | -106.7969
17.3875 | -106.7995
17.3911 | -106.8005
17.3948 | -106.7996
17.3984 | -106.7994
17.402 | -106.7988
17.4056 | -106.7987
17.4092 | -106.802
17.4128 | -106.8059
17.4164 | -106.8093
17.4201 | -106.8143
17.4237 | -106.8203
17.4273 | -106.8278
17.4309 | -106.8341
17.4345 | -106.8404
17.4381 | -106.8491
17.4417 | -106.8597
17.4453 | -106.8703
17.449 | -106.8816
17.4526 | -106.8943
17.4562 | -106.9062
17.4598 | -106.9156
17.4634 | -106.9259
17.467 | -106.939
17.4706 | -106.9508
17.4743 | -106.9628
17.4779 | -106.9753
17.4815 | -106.9876
17.4851 | -106.9966
17.4887 | -107.0051
17.4923 | -107.0118
17.4959 | -107.0174
17.4995 | -107.0216
17.5032 | -107.0268
17.5068 | -107.0329
17.5104 | -107.0368
17.514 | -107.0384
17.5176 | -107.0382
17.5212 | -107.0337
17.5248 | -107.0311
17.5285 | -107.0283
17.5321 | -107.0224
17.5357 | -107.0164
17.5393 | -107.0086
17.5429 | -106.9981
17.5465 | -106.9903
17.5501 | -106.9825
17.5537 | -106.9725
17.5574 | -106.9616
17.561 | -106.9508
17.5646 | -106.9381
17.5682 | -106.9249
17.5718 | -106.9105
17.5754 | -106.8982
17.579 | -106.8864
17.5827 | -106.8772
17.5863 | -106.8678
17.5899 | -106.8562
17.5935 | -106.8428
17.5971 | -106.8311
17.6007 | -106.8211
17.6043 | -106.81
17.6079 | -106.7976
17.6116 | -106.7876
17.6152 | -106.7816
17.6188 | -106.7769
17.6224 | -106.7737
17.626 | -106.7712
17.6296 | -106.7681
17.6332 | -106.7644
17.6369 | -106.7613
17.6405 | -106.7608
17.6441 | -106.7632
17.6477 | -106.7664
17.6513 | -106.7701
17.6549 | -106.772
17.6585 | -106.7732
17.6622 | -106.7753
17.6658 | -106.7795
17.6694 | -106.782
17.673 | -106.7862
17.6766 | -106.7922
17.6802 | -106.8002
17.6838 | -106.8095
17.6874 | -106.8174
17.6911 | -106.8256
17.6947 | -106.8339
17.6983 | -106.8436
17.7019 | -106.8517
17.7055 | -106.8628
17.7091 | -106.8759
17.7127 | -106.8875
17.7163 | -106.8974
17.72 | -106.909
17.7236 | -106.9205
17.7272 | -106.9317
17.7308 | -106.943
17.7344 | -106.9545
17.738 | -106.9673
17.7416 | -106.9788
17.7453 | -106.9911
17.7489 | -107.0043
17.7525 | -107.0152
17.7561 | -107.0221
17.7597 | -107.0309
17.7633 | -107.0446
17.7669 | -107.0578
17.7706 | -107.0673
17.7742 | -107.0734
17.7778 | -107.0753
17.7814 | -107.0776
17.785 | -107.0814
17.7886 | -107.085
17.7922 | -107.0872
17.7958 | -107.0881
17.7995 | -107.087
17.8031 | -107.0848
17.8067 | -107.0811
17.8103 | -107.0756
17.8139 | -107.069
17.8175 | -107.0632
17.8211 | -107.0538
17.8248 | -107.0429
17.8284 | -107.031
17.832 | -107.0169
17.8356 | -106.9978
17.8392 | -106.9819
17.8428 | -106.9671
17.8464 | -106.9497
17.85 | -106.9321
17.8537 | -106.9165
17.8573 | -106.9009
17.8609 | -106.8817
17.8645 | -106.8626
17.8681 | -106.8426
17.8717 | -106.8207
17.8753 | -106.7995
17.879 | -106.7798
17.8826 | -106.7588
17.8862 | -106.7378
17.8898 | -106.7167
17.8934 | -106.6954
17.897 | -106.6737
17.9006 | -106.6546
17.9042 | -106.6351
17.9079 | -106.6175
17.9115 | -106.6009
17.9151 | -106.585
17.9187 | -106.5691
17.9223 | -106.5545
17.9259 | -106.5407
17.9295 | -106.5255
17.9332 | -106.5123
17.9368 | -106.5003
17.9404 | -106.4901
17.944 | -106.4825
17.9476 | -106.4772
17.9512 | -106.473
17.9548 | -106.4701
17.9584 | -106.4661
17.9621 | -106.4623
17.9657 | -106.4637
17.9693 | -106.4683
17.9729 | -106.4736
17.9765 | -106.4797
17.9801 | -106.4873
17.9837 | -106.4964
17.9874 | -106.5063
17.991 | -106.5178
17.9946 | -106.5314
17.9982 | -106.5452
18.0018 | -106.5587
18.0054 | -106.5732
18.009 | -106.5905
18.0126 | -106.6104
18.0163 | -106.6279
18.0199 | -106.6438
18.0235 | -106.6602
18.0271 | -106.6775
18.0307 | -106.6989
18.0343 | -106.7202
18.0379 | -106.7376
18.0416 | -106.755
18.0452 | -106.7709
18.0488 | -106.7857
18.0524 | -106.8025
18.056 | -106.819
18.0596 | -106.8342
18.0632 | -106.8505
18.0668 | -106.8655
18.0705 | -106.8767
18.0741 | -106.8862
18.0777 | -106.8975
18.0813 | -106.9116
18.0849 | -106.9241
18.0885 | -106.9356
18.0921 | -106.9484
18.0958 | -106.9571
18.0994 | -106.9664
18.103 | -106.9753
18.1066 | -106.9817
18.1102 | -106.9858
18.1138 | -106.9927
18.1174 | -107.0023
18.121 | -107.01
18.1247 | -107.0162
18.1283 | -107.0226
18.1319 | -107.0278
18.1355 | -107.0316
18.1391 | -107.0375
18.1427 | -107.0426
18.1463 | -107.0456
18.15 | -107.049
18.1536 | -107.0511
18.1572 | -107.0502
18.1608 | -107.052
18.1644 | -107.0544
18.168 | -107.0561
18.1716 | -107.0553
18.1752 | -107.0553
18.1789 | -107.0552
18.1825 | -107.054
18.1861 | -107.0529
18.1897 | -107.0528
18.1933 | -107.053
18.1969 | -107.0532
18.2005 | -107.0508
18.2042 | -107.0493
18.2078 | -107.048
18.2114 | -107.0466
18.215 | -107.0451
18.2186 | -107.0426
18.2222 | -107.0364
18.2258 | -107.0306
18.2294 | -107.0285
18.2331 | -107.0258
18.2367 | -107.0216
18.2403 | -107.0158
18.2439 | -107.0095
18.2475 | -107.0049
18.2511 | -106.9981
18.2547 | -106.9893
18.2584 | -106.9824
18.262 | -106.977
18.2656 | -106.9701
18.2692 | -106.9626
18.2728 | -106.9561
18.2764 | -106.9505
18.28 | -106.9441
18.2837 | -106.9377
18.2873 | -106.9335
18.2909 | -106.9288
18.2945 | -106.9217
18.2981 | -106.9145
18.3017 | -106.9087
18.3053 | -106.9013
18.3089 | -106.8952
18.3126 | -106.8888
18.3162 | -106.8827
18.3198 | -106.8784
18.3234 | -106.8738
18.327 | -106.8687
18.3306 | -106.8663
18.3342 | -106.8654
18.3378 | -106.8643
18.3415 | -106.8599
18.3451 | -106.8555
18.3487 | -106.8532
18.3523 | -106.8511
18.3559 | -106.8499
18.3595 | -106.8485
18.3631 | -106.8485
18.3668 | -106.8502
18.3704 | -106.8511
18.374 | -106.8502
18.3776 | -106.8498
18.3812 | -106.8484
18.3848 | -106.8475
18.3884 | -106.8478
18.3921 | -106.8473
18.3957 | -106.8484
18.3993 | -106.8491
18.4029 | -106.8498
18.4065 | -106.852
18.4101 | -106.852
18.4137 | -106.8516
18.4173 | -106.8533
18.421 | -106.8518
18.4246 | -106.8489
18.4282 | -106.847
18.4318 | -106.8485
18.4354 | -106.8498
18.439 | -106.85
18.4426 | -106.8505
18.4463 | -106.8508
18.4499 | -106.8466
18.4535 | -106.8412
18.4571 | -106.8366
18.4607 | -106.8327
18.4643 | -106.8295
18.4679 | -106.827
18.4715 | -106.8242
18.4752 | -106.8212
18.4788 | -106.8164
18.4824 | -106.8107
18.486 | -106.8052
18.4896 | -106.7984
18.4932 | -106.7939
18.4968 | -106.7909
18.5005 | -106.7864
18.5041 | -106.7807
18.5077 | -106.7755
18.5113 | -106.7713
18.5149 | -106.7683
18.5185 | -106.7652
18.5221 | -106.7599
18.5257 | -106.7537
18.5294 | -106.7489
18.533 | -106.7435
18.5366 | -106.7383
18.5402 | -106.7343
18.5438 | -106.7313
18.5474 | -106.7302
18.551 | -106.7274
18.5547 | -106.7227
18.5583 | -106.7204
18.5619 | -106.7201
18.5655 | -106.7174
18.5691 | -106.7176
18.5727 | -106.7181
18.5763 | -106.7163
18.5799 | -106.7127
18.5836 | -106.7104
18.5872 | -106.7085
18.5908 | -106.7074
18.5944 | -106.7043
18.598 | -106.7006
18.6016 | -106.697
18.6052 | -106.6941
18.6089 | -106.691
18.6125 | -106.6885
18.6161 | -106.6868
18.6197 | -106.6849
18.6233 | -106.6826
18.6269 | -106.6808
18.6305 | -106.6779
18.6341 | -106.676
18.6378 | -106.6765
18.6414 | -106.6761
18.645 | -106.677
18.6486 | -106.6791
18.6522 | -106.6812
18.6558 | -106.6834
18.6594 | -106.6867
18.6631 | -106.6909
18.6667 | -106.6954
18.6703 | -106.6983
18.6739 | -106.7029
18.6775 | -106.7078
18.6811 | -106.7118
18.6847 | -106.7183
18.6883 | -106.7259
18.692 | -106.7329
18.6956 | -106.7378
18.6992 | -106.7433
18.7028 | -106.7508
18.7064 | -106.7574
18.71 | -106.7656
18.7136 | -106.7756
18.7173 | -106.7839
18.7209 | -106.7908
18.7245 | -106.7985
18.7281 | -106.8054
18.7317 | -106.8102
18.7353 | -106.8137
18.7389 | -106.8184
18.7425 | -106.8232
18.7462 | -106.8276
18.7498 | -106.8287
18.7534 | -106.8271
18.757 | -106.8242
18.7606 | -106.8228
18.7642 | -106.8187
18.7678 | -106.8183
18.7715 | -106.8177
18.7751 | -106.8154
18.7787 | -106.8105
18.7823 | -106.8035
18.7859 | -106.7993
18.7895 | -106.795
18.7931 | -106.7879
18.7967 | -106.7801
18.8004 | -106.7722
18.804 | -106.7634
18.8076 | -106.7543
18.8112 | -106.7447
18.8148 | -106.734
18.8184 | -106.7219
18.822 | -106.7137
18.8257 | -106.707
18.8293 | -106.697
18.8329 | -106.6865
18.8365 | -106.6743
18.8401 | -106.6616
18.8437 | -106.6528
18.8473 | -106.6457
18.8509 | -106.6376
18.8546 | -106.6288
18.8582 | -106.62
18.8618 | -106.614
18.8654 | -106.6095
18.869 | -106.6034
18.8726 | -106.5972
18.8762 | -106.5913
18.8799 | -106.5857
18.8835 | -106.5797
18.8871 | -106.5747
18.8907 | -106.5718
18.8943 | -106.5692
18.8979 | -106.5673
18.9015 | -106.5671
18.9051 | -106.5664
18.9088 | -106.5671
18.9124 | -106.5691
18.916 | -106.5712
18.9196 | -106.5734
18.9232 | -106.5753
18.9268 | -106.577
18.9304 | -106.581
18.9341 | -106.5857
18.9377 | -106.5931
18.9413 | -106.6008
18.9449 | -106.6083
18.9485 | -106.6185
18.9521 | -106.6274
18.9557 | -106.6343
18.9594 | -106.642
18.963 | -106.6527
18.9666 | -106.6628
18.9702 | -106.6708
18.9738 | -106.6785
18.9774 | -106.6865
18.981 | -106.6949
18.9846 | -106.7051
18.9883 | -106.7129
18.9919 | -106.7196
18.9955 | -106.7273
18.9991 | -106.7366
19.0027 | -106.7462
19.0063 | -106.7551
19.0099 | -106.7645
19.0135 | -106.7725
19.0172 | -106.7795
19.0208 | -106.787
19.0244 | -106.7921
19.028 | -106.7961
19.0316 | -106.7984
19.0352 | -106.7987
19.0388 | -106.7999
19.0425 | -106.8017
19.0461 | -106.8035
19.0497 | -106.8042
19.0533 | -106.807
19.0569 | -106.8069
19.0605 | -106.8085
19.0641 | -106.8115
19.0678 | -106.8123
19.0714 | -106.8131
19.075 | -106.8135
19.0786 | -106.8107
19.0822 | -106.807
19.0858 | -106.8026
19.0894 | -106.7996
19.093 | -106.7978
19.0967 | -106.7956
19.1003 | -106.7923
19.1039 | -106.7911
19.1075 | -106.7913
19.1111 | -106.79
19.1147 | -106.7853
19.1183 | -106.782
19.122 | -106.7781
19.1256 | -106.7764
19.1292 | -106.7762
19.1328 | -106.7732
19.1364 | -106.7692
19.14 | -106.7644
19.1436 | -106.7616
19.1472 | -106.761
19.1509 | -106.758
19.1545 | -106.7559
19.1581 | -106.754
19.1617 | -106.7504
19.1653 | -106.7488
19.1689 | -106.7469
19.1725 | -106.7424
19.1762 | -106.7384
19.1798 | -106.7345
19.1834 | -106.7321
19.187 | -106.73
19.1906 | -106.7282
19.1942 | -106.7268
19.1978 | -106.724
19.2014 | -106.721
19.2051 | -106.7176
19.2087 | -106.7147
19.2123 | -106.7117
19.2159 | -106.709
19.2195 | -106.7038
19.2231 | -106.701
19.2267 | -106.7
19.2304 | -106.698
19.234 | -106.6965
19.2376 | -106.6938
19.2412 | -106.6916
19.2448 | -106.6916
19.2484 | -106.6903
19.252 | -106.6894
19.2556 | -106.6884
19.2593 | -106.6886
19.2629 | -106.6884
19.2665 | -106.6863
19.2701 | -106.6848
19.2737 | -106.6828
19.2773 | -106.6799
19.2809 | -106.6781
19.2846 | -106.6779
19.2882 | -106.6762
19.2918 | -106.6748
19.2954 | -106.6732
19.299 | -106.6723
19.3026 | -106.6704
19.3062 | -106.6686
19.3098 | -106.6665
19.3135 | -106.6657
19.3171 | -106.6641
19.3207 | -106.6621
19.3243 | -106.6606
19.3279 | -106.6561
19.3315 | -106.6526
19.3351 | -106.6478
19.3388 | -106.6439
19.3424 | -106.6403
19.346 | -106.6337
19.3496 | -106.6283
19.3532 | -106.6261
19.3568 | -106.6244
19.3604 | -106.6227
19.364 | -106.6204
19.3677 | -106.6123
19.3713 | -106.6058
19.3749 | -106.603
19.3785 | -106.598
19.3821 | -106.592
19.3857 | -106.5875
19.3893 | -106.5842
19.393 | -106.5786
19.3966 | -106.5728
19.4002 | -106.5667
19.4038 | -106.5598
19.4074 | -106.5557
19.411 | -106.5522
19.4146 | -106.5489
19.4182 | -106.5449
19.4219 | -106.5395
19.4255 | -106.534
19.4291 | -106.5272
19.4327 | -106.5203
19.4363 | -106.5157
19.4399 | -106.5105
19.4435 | -106.5055
19.4472 | -106.5042
19.4508 | -106.5011
19.4544 | -106.4979
19.458 | -106.4945
19.4616 | -106.4919
19.4652 | -106.4899
19.4688 | -106.4891
19.4724 | -106.4889
19.4761 | -106.4901
19.4797 | -106.4936
19.4833 | -106.4967
19.4869 | -106.5004
19.4905 | -106.5028
19.4941 | -106.5069
19.4977 | -106.5099
19.5014 | -106.514
19.505 | -106.5195
19.5086 | -106.5242
19.5122 | -106.5313
19.5158 | -106.5394
19.5194 | -106.547
19.523 | -106.5544
19.5266 | -106.5598
19.5303 | -106.5674
19.5339 | -106.5747
19.5375 | -106.5817
19.5411 | -106.59
19.5447 | -106.5974
19.5483 | -106.6051
19.5519 | -106.6119
19.5556 | -106.618
19.5592 | -106.6239
19.5628 | -106.6292
19.5664 | -106.6334
19.57 | -106.6392
19.5736 | -106.6432
19.5772 | -106.6477
19.5808 | -106.6517
19.5845 | -106.6547
19.5881 | -106.6585
19.5917 | -106.6609
19.5953 | -106.664
19.5989 | -106.6669
19.6025 | -106.6693
19.6061 | -106.6728
19.6098 | -106.6751
19.6134 | -106.6777
19.617 | -106.6796
19.6206 | -106.6833
19.6242 | -106.6883
19.6278 | -106.6887
19.6314 | -106.6869
19.635 | -106.6871
19.6387 | -106.6878
19.6423 | -106.6874
19.6459 | -106.6897
19.6495 | -106.6914
19.6531 | -106.6911
19.6567 | -106.6899
19.6603 | -106.6893
19.664 | -106.6898
19.6676 | -106.6901
19.6712 | -106.6903
19.6748 | -106.6923
19.6784 | -106.6958
19.682 | -106.6983
19.6856 | -106.6974
19.6893 | -106.6995
19.6929 | -106.7032
19.6965 | -106.7057
19.7001 | -106.7081
19.7037 | -106.7117
19.7073 | -106.7166
19.7109 | -106.7213
19.7145 | -106.7249
19.7182 | -106.7289
19.7218 | -106.7371
19.7254 | -106.7421
19.729 | -106.7472
19.7326 | -106.7538
19.7362 | -106.7608
19.7398 | -106.7664
19.7435 | -106.7746
19.7471 | -106.7847
19.7507 | -106.793
19.7543 | -106.7979
19.7579 | -106.8027
19.7615 | -106.8093
19.7651 | -106.8162
19.7687 | -106.8241
19.7724 | -106.8309
19.776 | -106.8371
19.7796 | -106.841
19.7832 | -106.8447
19.7868 | -106.8451
19.7904 | -106.8445
19.794 | -106.8449
19.7977 | -106.8452
19.8013 | -106.8465
19.8049 | -106.8445
19.8085 | -106.8365
19.8121 | -106.8318
19.8157 | -106.827
19.8193 | -106.822
19.8229 | -106.8168
19.8266 | -106.8102
19.8302 | -106.8049
19.8338 | -106.8005
19.8374 | -106.7957
19.841 | -106.7898
19.8446 | -106.7844
19.8482 | -106.7804
19.8519 | -106.7734
19.8555 | -106.766
19.8591 | -106.7568
19.8627 | -106.7497
19.8663 | -106.7422
19.8699 | -106.7352
19.8735 | -106.7284
19.8771 | -106.7198
19.8808 | -106.7134
19.8844 | -106.7098
19.888 | -106.7056
19.8916 | -106.7017
19.8952 | -106.6996
19.8988 | -106.6968
19.9024 | -106.6915
19.9061 | -106.6878
19.9097 | -106.6872
19.9133 | -106.687
19.9169 | -106.6854
19.9205 | -106.686
19.9241 | -106.6877
19.9277 | -106.6879
19.9313 | -106.6861
19.935 | -106.6869
19.9386 | -106.6905
19.9422 | -106.6929
19.9458 | -106.6941
19.9494 | -106.6954
19.953 | -106.6975
19.9566 | -106.6994
19.9603 | -106.7016
19.9639 | -106.7057
19.9675 | -106.7121
19.9711 | -106.7155
19.9747 | -106.719
19.9783 | -106.7253
19.9819 | -106.7315
19.9855 | -106.7376
19.9892 | -106.7436
19.9928 | -106.7492
19.9964 | -106.7553
20 | -106.7605
20.0036 | -106.7644
20.0072 | -106.7706
20.0108 | -106.778
20.0145 | -106.7854
20.0181 | -106.7923
20.0217 | -106.7997
20.0253 | -106.8027
20.0289 | -106.8049
20.0325 | -106.8079
20.0361 | -106.8137
20.0397 | -106.8202
20.0434 | -106.8256
20.047 | -106.8312
20.0506 | -106.8349
20.0542 | -106.8382
20.0578 | -106.8438
20.0614 | -106.8503
20.065 | -106.8566
20.0687 | -106.8624
20.0723 | -106.8659
20.0759 | -106.8669
20.0795 | -106.8672
20.0831 | -106.8659
20.0867 | -106.8677
20.0903 | -106.8661
20.0939 | -106.8658
20.0976 | -106.8655
20.1012 | -106.8639
20.1048 | -106.8609
20.1084 | -106.8566
20.112 | -106.8541
20.1156 | -106.85
20.1192 | -106.8468
20.1229 | -106.8447
20.1265 | -106.843
20.1301 | -106.8402
20.1337 | -106.8391
20.1373 | -106.8375
20.1409 | -106.8373
20.1445 | -106.8376
20.1481 | -106.8352
20.1518 | -106.835
20.1554 | -106.8346
20.159 | -106.8329
20.1626 | -106.8346
20.1662 | -106.8352
20.1698 | -106.8358
20.1734 | -106.8348
20.1771 | -106.8349
20.1807 | -106.8346
20.1843 | -106.8353
20.1879 | -106.8394
20.1915 | -106.844
20.1951 | -106.8478
20.1987 | -106.8521
20.2023 | -106.8564
20.206 | -106.859
20.2096 | -106.8609
20.2132 | -106.8632
20.2168 | -106.8673
20.2204 | -106.8692
20.224 | -106.8701
20.2276 | -106.8734
20.2313 | -106.8766
20.2349 | -106.8794
20.2385 | -106.8826
20.2421 | -106.8838
20.2457 | -106.8842
20.2493 | -106.8857
20.2529 | -106.8864
20.2565 | -106.8876
20.2602 | -106.8881
20.2638 | -106.8851
20.2674 | -106.8844
20.271 | -106.8837
20.2746 | -106.8828
20.2782 | -106.8814
20.2818 | -106.8801
20.2855 | -106.8776
20.2891 | -106.8741
20.2927 | -106.8684
20.2963 | -106.8616
20.2999 | -106.8551
20.3035 | -106.8481
20.3071 | -106.8424
20.3107 | -106.8397
20.3144 | -106.8368
20.318 | -106.8336
20.3216 | -106.8274
20.3252 | -106.8191
20.3288 | -106.8089
20.3324 | -106.8006
20.336 | -106.7931
20.3397 | -106.7892
20.3433 | -106.7838
20.3469 | -106.7753
20.3505 | -106.7674
20.3541 | -106.7612
20.3577 | -106.7567
20.3613 | -106.7541
20.365 | -106.7521
20.3686 | -106.7493
20.3722 | -106.7471
20.3758 | -106.7458
20.3794 | -106.7455
20.383 | -106.7426
20.3866 | -106.7394
20.3902 | -106.7365
20.3939 | -106.7343
20.3975 | -106.7337
20.4011 | -106.7336
20.4047 | -106.7321
20.4083 | -106.7317
20.4119 | -106.7299
20.4155 | -106.7298
20.4192 | -106.7282
20.4228 | -106.725
20.4264 | -106.7247
20.43 | -106.7263
20.4336 | -106.727
20.4372 | -106.7273
20.4408 | -106.7272
20.4444 | -106.7267
20.4481 | -106.7287
20.4517 | -106.7287
20.4553 | -106.7281
20.4589 | -106.7279
20.4625 | -106.7309
20.4661 | -106.7323
20.4697 | -106.7335
20.4734 | -106.7347
20.477 | -106.7336
20.4806 | -106.7314
20.4842 | -106.7299
20.4878 | -106.727
20.4914 | -106.7229
20.495 | -106.7213
20.4986 | -106.7199
20.5023 | -106.7158
20.5059 | -106.7121
20.5095 | -106.7091
20.5131 | -106.7055
20.5167 | -106.6999
20.5203 | -106.6942
20.5239 | -106.6881
20.5276 | -106.6798
20.5312 | -106.6721
20.5348 | -106.6637
20.5384 | -106.6538
20.542 | -106.6454
20.5456 | -106.6372
20.5492 | -106.6301
20.5528 | -106.6235
20.5565 | -106.6157
20.5601 | -106.6078
20.5637 | -106.599
20.5673 | -106.5896
20.5709 | -106.58
20.5745 | -106.5729
20.5781 | -106.5659
20.5818 | -106.56
20.5854 | -106.5543
20.589 | -106.5469
20.5926 | -106.5413
20.5962 | -106.5355
20.5998 | -106.5295
20.6034 | -106.5243
20.607 | -106.5191
20.6107 | -106.515
20.6143 | -106.5122
20.6179 | -106.511
20.6215 | -106.5094
20.6251 | -106.5064
20.6287 | -106.5046
20.6323 | -106.5023
20.636 | -106.5
20.6396 | -106.4994
20.6432 | -106.499
20.6468 | -106.4978
20.6504 | -106.498
20.654 | -106.4999
20.6576 | -106.5011
20.6612 | -106.5016
20.6649 | -106.5043
20.6685 | -106.5066
20.6721 | -106.5093
20.6757 | -106.5088
20.6793 | -106.5111
20.6829 | -106.5161
20.6865 | -106.5208
20.6902 | -106.5226
20.6938 | -106.5251
20.6974 | -106.5287
20.701 | -106.5299
20.7046 | -106.5317
20.7082 | -106.5333
20.7118 | -106.5371
20.7154 | -106.5409
20.7191 | -106.5435
20.7227 | -106.5444
20.7263 | -106.5459
20.7299 | -106.5476
20.7335 | -106.549
20.7371 | -106.5489
20.7407 | -106.5484
20.7444 | -106.5483
20.748 | -106.5476
20.7516 | -106.5464
20.7552 | -106.5447
20.7588 | -106.5439
20.7624 | -106.543
20.766 | -106.5435
20.7696 | -106.5462
20.7733 | -106.548
20.7769 | -106.5498
20.7805 | -106.5538
20.7841 | -106.5578
20.7877 | -106.5589
20.7913 | -106.5596
20.7949 | -106.5645
20.7986 | -106.569
20.8022 | -106.5719
20.8058 | -106.577
20.8094 | -106.5831
20.813 | -106.5898
20.8166 | -106.5975
20.8202 | -106.6054
20.8238 | -106.612
20.8275 | -106.6178
20.8311 | -106.6246
20.8347 | -106.6304
20.8383 | -106.6366
20.8419 | -106.6422
20.8455 | -106.6486
20.8491 | -106.655
20.8528 | -106.6611
20.8564 | -106.6671
20.86 | -106.6735
20.8636 | -106.6798
20.8672 | -106.6861
20.8708 | -106.6913
20.8744 | -106.6966
20.878 | -106.7002
20.8817 | -106.7042
20.8853 | -106.7083
20.8889 | -106.7088
20.8925 | -106.7084
20.8961 | -106.7096
20.8997 | -106.7095
20.9033 | -106.7105
20.907 | -106.7107
20.9106 | -106.7107
20.9142 | -106.7084
20.9178 | -106.7074
20.9214 | -106.7066
20.925 | -106.7048
20.9286 | -106.703
20.9322 | -106.7013
20.9359 | -106.6997
20.9395 | -106.6982
20.9431 | -106.6945
20.9467 | -106.6891
20.9503 | -106.6846
20.9539 | -106.6827
20.9575 | -106.6837
20.9612 | -106.6835
20.9648 | -106.6835
20.9684 | -106.6844
20.972 | -106.6886
20.9756 | -106.6903
20.9792 | -106.6923
20.9828 | -106.6945
20.9865 | -106.6991
20.9901 | -106.7038
20.9937 | -106.7093
20.9973 | -106.7151
21.0009 | -106.7213
21.0045 | -106.7276
21.0081 | -106.7345
21.0117 | -106.7423
21.0154 | -106.7528
21.019 | -106.7632
21.0226 | -106.774
21.0262 | -106.7853
21.0298 | -106.7972
21.0334 | -106.8113
21.037 | -106.8251
21.0406 | -106.8403
21.0443 | -106.8555
21.0479 | -106.8701
21.0515 | -106.8839
21.0551 | -106.897
21.0587 | -106.9104
21.0623 | -106.9239
21.0659 | -106.9378
21.0696 | -106.9524
21.0732 | -106.967
21.0768 | -106.979
21.0804 | -106.9879
21.084 | -106.9995
21.0876 | -107.0131
21.0912 | -107.0241
21.0949 | -107.0341
21.0985 | -107.0446
21.1021 | -107.0539
21.1057 | -107.0625
21.1093 | -107.0681
21.1129 | -107.0724
21.1165 | -107.0776
21.1201 | -107.0854
21.1238 | -107.0921
21.1274 | -107.0947
21.131 | -107.0957
21.1346 | -107.0967
21.1382 | -107.0967
21.1418 | -107.0938
21.1454 | -107.0906
21.1491 | -107.0858
21.1527 | -107.0801
21.1563 | -107.0745
21.1599 | -107.0685
21.1635 | -107.0631
21.1671 | -107.0562
21.1707 | -107.0493
21.1743 | -107.0407
21.178 | -107.0347
21.1816 | -107.0278
21.1852 | -107.0194
21.1888 | -107.0122
21.1924 | -107.0037
21.196 | -106.9938
21.1996 | -106.9845
21.2033 | -106.9763
21.2069 | -106.9699
21.2105 | -106.9608
21.2141 | -106.9549
21.2177 | -106.9488
21.2213 | -106.9404
21.2249 | -106.9324
21.2285 | -106.9251
21.2322 | -106.9187
21.2358 | -106.9134
21.2394 | -106.9076
21.243 | -106.9046
21.2466 | -106.9018
21.2502 | -106.8976
21.2538 | -106.8933
21.2575 | -106.8886
21.2611 | -106.8843
21.2647 | -106.8828
21.2683 | -106.8825
21.2719 | -106.8827
21.2755 | -106.8846
21.2791 | -106.8853
21.2827 | -106.8857
21.2864 | -106.8855
21.29 | -106.8873
21.2936 | -106.8883
21.2972 | -106.8897
21.3008 | -106.8933
21.3044 | -106.8986
21.308 | -106.9033
21.3117 | -106.9067
21.3153 | -106.9087
21.3189 | -106.9114
21.3225 | -106.9155
21.3261 | -106.9186
21.3297 | -106.9205
21.3333 | -106.9228
21.3369 | -106.9239
21.3406 | -106.9255
21.3442 | -106.9283
21.3478 | -106.932
21.3514 | -106.9371
21.355 | -106.9406
21.3586 | -106.944
21.3622 | -106.9464
21.3659 | -106.9485
21.3695 | -106.951
21.3731 | -106.9548
21.3767 | -106.957
21.3803 | -106.9607
21.3839 | -106.9673
21.3875 | -106.9712
21.3911 | -106.9736
21.3948 | -106.9756
21.3984 | -106.9755
21.402 | -106.9776
21.4056 | -106.9795
21.4092 | -106.9807
21.4128 | -106.9829
21.4164 | -106.9856
21.4201 | -106.9887
21.4237 | -106.9891
21.4273 | -106.9899
21.4309 | -106.9907
21.4345 | -106.9922
21.4381 | -106.9911
21.4417 | -106.9906
21.4453 | -106.9915
21.449 | -106.9891
21.4526 | -106.9877
21.4562 | -106.9878
21.4598 | -106.9869
21.4634 | -106.9854
21.467 | -106.9814
21.4706 | -106.9779
21.4743 | -106.9732
21.4779 | -106.9708
21.4815 | -106.9662
21.4851 | -106.9595
21.4887 | -106.9531
21.4923 | -106.947
21.4959 | -106.9431
21.4995 | -106.9387
21.5032 | -106.9332
21.5068 | -106.9267
21.5104 | -106.9196
21.514 | -106.9135
21.5176 | -106.9035
21.5212 | -106.8934
21.5248 | -106.8864
21.5285 | -106.8786
21.5321 | -106.8684
21.5357 | -106.8582
21.5393 | -106.8486
21.5429 | -106.8389
21.5465 | -106.8291
21.5501 | -106.8192
21.5537 | -106.8084
21.5574 | -106.7968
21.561 | -106.7835
21.5646 | -106.7709
21.5682 | -106.7572
21.5718 | -106.7446
21.5754 | -106.7337
21.579 | -106.7224
21.5827 | -106.7114
21.5863 | -106.7027
21.5899 | -106.6937
21.5935 | -106.6826
21.5971 | -106.6747
21.6007 | -106.6682
21.6043 | -106.6594
21.6079 | -106.6502
21.6116 | -106.6414
21.6152 | -106.6336
21.6188 | -106.6243
21.6224 | -106.6171
21.626 | -106.6099
21.6296 | -106.6007
21.6332 | -106.5935
21.6369 | -106.5869
21.6405 | -106.5818
21.6441 | -106.5768
21.6477 | -106.5714
21.6513 | -106.5636
21.6549 | -106.5563
21.6585 | -106.5515
21.6622 | -106.5462
21.6658 | -106.5419
21.6694 | -106.5367
21.673 | -106.5306
21.6766 | -106.5248
21.6802 | -106.5197
21.6838 | -106.5162
21.6874 | -106.5123
21.6911 | -106.5097
21.6947 | -106.5072
21.6983 | -106.5052
21.7019 | -106.5025
21.7055 | -106.4994
21.7091 | -106.4955
21.7127 | -106.4924
21.7163 | -106.489
21.72 | -106.4881
21.7236 | -106.4871
21.7272 | -106.4841
21.7308 | -106.4813
21.7344 | -106.4788
21.738 | -106.4756
21.7416 | -106.4711
21.7453 | -106.4689
21.7489 | -106.4656
21.7525 | -106.4625
21.7561 | -106.4601
21.7597 | -106.4555
21.7633 | -106.4497
21.7669 | -106.4445
21.7706 | -106.4393
21.7742 | -106.4354
21.7778 | -106.4303
21.7814 | -106.424
21.785 | -106.4185
21.7886 | -106.4136
21.7922 | -106.4109
21.7958 | -106.4077
21.7995 | -106.4036
21.8031 | -106.3978
21.8067 | -106.3921
21.8103 | -106.3872
21.8139 | -106.3827
21.8175 | -106.3781
21.8211 | -106.375
21.8248 | -106.3742
21.8284 | -106.3713
21.832 | -106.3658
21.8356 | -106.3616
21.8392 | -106.3582
21.8428 | -106.3546
21.8464 | -106.3493
21.85 | -106.3447
21.8537 | -106.3409
21.8573 | -106.3391
21.8609 | -106.3384
21.8645 | -106.3387
21.8681 | -106.3387
21.8717 | -106.338
21.8753 | -106.338
21.879 | -106.3361
21.8826 | -106.3334
21.8862 | -106.3306
21.8898 | -106.3287
21.8934 | -106.328
21.897 | -106.3272
21.9006 | -106.3257
21.9042 | -106.3257
21.9079 | -106.3241
21.9115 | -106.3198
21.9151 | -106.3179
21.9187 | -106.3165
21.9223 | -106.3156
21.9259 | -106.3151
21.9295 | -106.3153
21.9332 | -106.315
21.9368 | -106.3152
21.9404 | -106.3167
21.944 | -106.3154
21.9476 | -106.3107
21.9512 | -106.3068
21.9548 | -106.3035
21.9584 | -106.3011
21.9621 | -106.3013
21.9657 | -106.3003
21.9693 | -106.297
21.9729 | -106.2943
21.9765 | -106.2916
21.9801 | -106.2906
21.9837 | -106.291
21.9874 | -106.2892
21.991 | -106.2876
21.9946 | -106.2868
21.9982 | -106.2858
22.0018 | -106.2844
22.0054 | -106.2834
22.009 | -106.2815
22.0126 | -106.2807
22.0163 | -106.2805
22.0199 | -106.2795
22.0235 | -106.2779
22.0271 | -106.278
22.0307 | -106.277
22.0343 | -106.2758
22.0379 | -106.2735
22.0416 | -106.2717
22.0452 | -106.271
22.0488 | -106.2737
22.0524 | -106.2761
22.056 | -106.2751
22.0596 | -106.2757
22.0632 | -106.2747
22.0668 | -106.2745
22.0705 | -106.2753
22.0741 | -106.2766
22.0777 | -106.2766
22.0813 | -106.2767
22.0849 | -106.2759
22.0885 | -106.2749
22.0921 | -106.2751
22.0958 | -106.2743
22.0994 | -106.2737
22.103 | -106.2725
22.1066 | -106.2729
22.1102 | -106.2727
22.1138 | -106.2705
22.1174 | -106.2686
22.121 | -106.2684
22.1247 | -106.2674
22.1283 | -106.2669
22.1319 | -106.2654
22.1355 | -106.2634
22.1391 | -106.2622
22.1427 | -106.2636
22.1463 | -106.2644
22.15 | -106.2665
22.1536 | -106.2686
22.1572 | -106.2709
22.1608 | -106.2726
22.1644 | -106.2727
22.168 | -106.2728
22.1716 | -106.2728
22.1752 | -106.2715
22.1789 | -106.2707
22.1825 | -106.2712
22.1861 | -106.272
22.1897 | -106.2744
22.1933 | -106.278
22.1969 | -106.2823
22.2005 | -106.2865
22.2042 | -106.2923
22.2078 | -106.297
22.2114 | -106.3011
22.215 | -106.3072
22.2186 | -106.3122
22.2222 | -106.319
22.2258 | -106.3289
22.2294 | -106.3366
22.2331 | -106.3438
22.2367 | -106.3505
22.2403 | -106.3583
22.2439 | -106.3681
22.2475 | -106.3771
22.2511 | -106.3854
22.2547 | -106.3956
22.2584 | -106.4073
22.262 | -106.422
22.2656 | -106.4354
22.2692 | -106.4479
22.2728 | -106.4603
22.2764 | -106.4719
22.28 | -106.4852
22.2837 | -106.4984
22.2873 | -106.5117
22.2909 | -106.5249
22.2945 | -106.5354
22.2981 | -106.5481
22.3017 | -106.5632
22.3053 | -106.5765
22.3089 | -106.5873
22.3126 | -106.5997
22.3162 | -106.6128
22.3198 | -106.6269
22.3234 | -106.6398
22.327 | -106.6514
22.3306 | -106.6645
22.3342 | -106.6767
22.3378 | -106.6864
22.3415 | -106.697
22.3451 | -106.7071
22.3487 | -106.7139
22.3523 | -106.7226
22.3559 | -106.7309
22.3595 | -106.7383
22.3631 | -106.7432
22.3668 | -106.7482
22.3704 | -106.7539
22.374 | -106.7554
22.3776 | -106.7544
22.3812 | -106.7541
22.3848 | -106.7543
22.3884 | -106.7518
22.3921 | -106.7492
22.3957 | -106.7461
22.3993 | -106.7415
22.4029 | -106.736
22.4065 | -106.7293
22.4101 | -106.7199
22.4137 | -106.7043
22.4173 | -106.6944
22.421 | -106.6813
22.4246 | -106.6671
22.4282 | -106.6569
22.4318 | -106.6479
22.4354 | -106.6358
22.439 | -106.6207
22.4426 | -106.6058
22.4463 | -106.5908
22.4499 | -106.5743
22.4535 | -106.5589
22.4571 | -106.5448
22.4607 | -106.5283
22.4643 | -106.512
22.4679 | -106.4948
22.4715 | -106.476
22.4752 | -106.4586
22.4788 | -106.4428
22.4824 | -106.4281
22.486 | -106.4118
22.4896 | -106.3975
22.4932 | -106.3849
22.4968 | -106.3742
22.5005 | -106.3622
22.5041 | -106.3496
22.5077 | -106.3372
22.5113 | -106.3258
22.5149 | -106.3152
22.5185 | -106.3047
22.5221 | -106.2939
22.5257 | -106.2841
22.5294 | -106.2762
22.533 | -106.2686
22.5366 | -106.2619
22.5402 | -106.254
22.5438 | -106.2464
22.5474 | -106.2404
22.551 | -106.2365
22.5547 | -106.2333
22.5583 | -106.2291
22.5619 | -106.2243
22.5655 | -106.2191
22.5691 | -106.2155
22.5727 | -106.2137
22.5763 | -106.2125
22.5799 | -106.2107
22.5836 | -106.2087
22.5872 | -106.2075
22.5908 | -106.208
22.5944 | -106.2063
22.598 | -106.2041
22.6016 | -106.2052
22.6052 | -106.2061
22.6089 | -106.2088
22.6125 | -106.2133
22.6161 | -106.2136
22.6197 | -106.2141
22.6233 | -106.2166
22.6269 | -106.2206
22.6305 | -106.2225
22.6341 | -106.2223
22.6378 | -106.2226
22.6414 | -106.2243
22.645 | -106.2255
22.6486 | -106.2257
22.6522 | -106.2247
22.6558 | -106.2247
22.6594 | -106.2225
22.6631 | -106.221
22.6667 | -106.2227
22.6703 | -106.2225
22.6739 | -106.221
22.6775 | -106.2173
22.6811 | -106.2124
22.6847 | -106.2074
22.6883 | -106.2037
22.692 | -106.2016
22.6956 | -106.2009
22.6992 | -106.1977
22.7028 | -106.1926
22.7064 | -106.1877
22.71 | -106.1814
22.7136 | -106.1765
22.7173 | -106.1719
22.7209 | -106.1661
22.7245 | -106.161
22.7281 | -106.1548
22.7317 | -106.1483
22.7353 | -106.1419
22.7389 | -106.1358
22.7425 | -106.1316
22.7462 | -106.1258
22.7498 | -106.1153
22.7534 | -106.1064
22.757 | -106.0992
22.7606 | -106.0922
22.7642 | -106.0848
22.7678 | -106.0773
22.7715 | -106.0716
22.7751 | -106.0636
22.7787 | -106.0561
22.7823 | -106.0492
22.7859 | -106.0414
22.7895 | -106.0334
22.7931 | -106.0264
22.7967 | -106.0166
22.8004 | -106.008
22.804 | -106.0005
22.8076 | -105.9945
22.8112 | -105.9872
22.8148 | -105.9812
22.8184 | -105.9758
22.822 | -105.9705
22.8257 | -105.9659
22.8293 | -105.9613
22.8329 | -105.9562
22.8365 | -105.9514
22.8401 | -105.9457
22.8437 | -105.9395
22.8473 | -105.934
22.8509 | -105.9298
22.8546 | -105.9266
22.8582 | -105.924
22.8618 | -105.9225
22.8654 | -105.9223
22.869 | -105.9213
22.8726 | -105.9182
22.8762 | -105.9167
22.8799 | -105.9169
22.8835 | -105.9154
22.8871 | -105.915
22.8907 | -105.9141
22.8943 | -105.9111
22.8979 | -105.9083
22.9015 | -105.9071
22.9051 | -105.9075
22.9088 | -105.9077
22.9124 | -105.9057
22.916 | -105.9064
22.9196 | -105.9063
22.9232 | -105.9048
22.9268 | -105.9039
22.9304 | -105.901
22.9341 | -105.9007
22.9377 | -105.9005
22.9413 | -105.9033
22.9449 | -105.9047
22.9485 | -105.9058
22.9521 | -105.9071
22.9557 | -105.9095
22.9594 | -105.9116
22.963 | -105.9116
22.9666 | -105.913
22.9702 | -105.916
22.9738 | -105.9195
22.9774 | -105.9241
22.981 | -105.9285
22.9846 | -105.9318
22.9883 | -105.9346
22.9919 | -105.9381
22.9955 | -105.9435
22.9991 | -105.9484
23.0027 | -105.9543
23.0063 | -105.9609
23.0099 | -105.9673
23.0135 | -105.9738
23.0172 | -105.9813
23.0208 | -105.9871
23.0244 | -105.9936
23.028 | -106.0012
23.0316 | -106.0095
23.0352 | -106.0193
23.0388 | -106.0258
23.0425 | -106.0316
23.0461 | -106.04
23.0497 | -106.0479
23.0533 | -106.0578
23.0569 | -106.0676
23.0605 | -106.0769
23.0641 | -106.0859
23.0678 | -106.0955
23.0714 | -106.1041
23.075 | -106.1139
23.0786 | -106.1234
23.0822 | -106.132
23.0858 | -106.1398
23.0894 | -106.1482
23.093 | -106.1559
23.0967 | -106.1638
23.1003 | -106.1725
23.1039 | -106.1794
23.1075 | -106.1866
23.1111 | -106.1937
23.1147 | -106.1993
23.1183 | -106.2051
23.122 | -106.2124
23.1256 | -106.2173
23.1292 | -106.2205
23.1328 | -106.2235
23.1364 | -106.2268
23.14 | -106.2295
23.1436 | -106.231
23.1472 | -106.2326
23.1509 | -106.2337
23.1545 | -106.2349
23.1581 | -106.2366
23.1617 | -106.238
23.1653 | -106.2403
23.1689 | -106.2426
23.1725 | -106.243
23.1762 | -106.2432
23.1798 | -106.2422
23.1834 | -106.2408
23.187 | -106.2394
23.1906 | -106.2374
23.1942 | -106.2352
23.1978 | -106.2343
23.2014 | -106.2342
23.2051 | -106.2331
23.2087 | -106.233
23.2123 | -106.2324
23.2159 | -106.2318
23.2195 | -106.2336
23.2231 | -106.2356
23.2267 | -106.2375
23.2304 | -106.2374
23.234 | -106.2363
23.2376 | -106.2357
23.2412 | -106.2364
23.2448 | -106.238
23.2484 | -106.2383
23.252 | -106.2385
23.2556 | -106.2411
23.2593 | -106.2428
23.2629 | -106.2426
23.2665 | -106.2405
23.2701 | -106.2387
23.2737 | -106.24
23.2773 | -106.2397
23.2809 | -106.237
23.2846 | -106.2371
23.2882 | -106.2392
23.2918 | -106.2424
23.2954 | -106.2441
23.299 | -106.2432
23.3026 | -106.2422
23.3062 | -106.2434
23.3098 | -106.2453
23.3135 | -106.2485
23.3171 | -106.2494
23.3207 | -106.2491
23.3243 | -106.2492
23.3279 | -106.2507
23.3315 | -106.2511
23.3351 | -106.2526
23.3388 | -106.2556
23.3424 | -106.2568
23.346 | -106.2573
23.3496 | -106.2606
23.3532 | -106.2647
23.3568 | -106.268
23.3604 | -106.2709
23.364 | -106.274
23.3677 | -106.2784
23.3713 | -106.2834
23.3749 | -106.287
23.3785 | -106.2904
23.3821 | -106.294
23.3857 | -106.2971
23.3893 | -106.2992
23.393 | -106.3029
23.3966 | -106.307
23.4002 | -106.3116
23.4038 | -106.3162
23.4074 | -106.3223
23.411 | -106.3293
23.4146 | -106.3373
23.4182 | -106.3458
23.4219 | -106.3534
23.4255 | -106.3584
23.4291 | -106.363
23.4327 | -106.3696
23.4363 | -106.3774
23.4399 | -106.3845
23.4435 | -106.3927
23.4472 | -106.3996
23.4508 | -106.4068
23.4544 | -106.4149
23.458 | -106.4244
23.4616 | -106.4332
23.4652 | -106.4402
23.4688 | -106.4478
23.4724 | -106.4563
23.4761 | -106.4658
23.4797 | -106.4745
23.4833 | -106.4834
23.4869 | -106.4935
23.4905 | -106.5027
23.4941 | -106.5111
23.4977 | -106.5204
23.5014 | -106.53
23.505 | -106.5391
23.5086 | -106.5471
23.5122 | -106.5547
23.5158 | -106.5614
23.5194 | -106.569
23.523 | -106.5763
23.5266 | -106.5839
23.5303 | -106.5898
23.5339 | -106.5954
23.5375 | -106.6013
23.5411 | -106.6065
23.5447 | -106.6112
23.5483 | -106.6169
23.5519 | -106.6209
23.5556 | -106.6231
23.5592 | -106.6264
23.5628 | -106.631
23.5664 | -106.635
23.57 | -106.6411
23.5736 | -106.6447
23.5772 | -106.6476
23.5808 | -106.6526
23.5845 | -106.6577
23.5881 | -106.6606
23.5917 | -106.6641
23.5953 | -106.669
23.5989 | -106.6729
23.6025 | -106.6782
23.6061 | -106.6822
23.6098 | -106.6876
23.6134 | -106.6946
23.617 | -106.6986
23.6206 | -106.704
23.6242 | -106.7096
23.6278 | -106.716
23.6314 | -106.7256
23.635 | -106.7341
23.6387 | -106.7391
23.6423 | -106.7448
23.6459 | -106.7504
23.6495 | -106.7605
23.6531 | -106.7713
23.6567 | -106.7813
23.6603 | -106.7912
23.664 | -106.8019
23.6676 | -106.8111
23.6712 | -106.8209
23.6748 | -106.8318
23.6784 | -106.8426
23.682 | -106.8534
23.6856 | -106.8633
23.6893 | -106.8736
23.6929 | -106.8824
23.6965 | -106.8922
23.7001 | -106.9044
23.7037 | -106.9171
23.7073 | -106.9294
23.7109 | -106.9405
23.7145 | -106.9508
23.7182 | -106.9597
23.7218 | -106.9694
23.7254 | -106.9792
23.729 | -106.9883
23.7326 | -106.9966
23.7362 | -107.0048
23.7398 | -107.0095
23.7435 | -107.0141
23.7471 | -107.0189
23.7507 | -107.023
23.7543 | -107.029
23.7579 | -107.0367
23.7615 | -107.0438
23.7651 | -107.0483
23.7687 | -107.0511
23.7724 | -107.0547
23.776 | -107.0583
23.7796 | -107.0631
23.7832 | -107.0657
23.7868 | -107.0672
23.7904 | -107.0696
23.794 | -107.0713
23.7977 | -107.0719
23.8013 | -107.0742
23.8049 | -107.0756
23.8085 | -107.077
23.8121 | -107.0783
23.8157 | -107.0809
23.8193 | -107.0813
23.8229 | -107.0805
23.8266 | -107.0788
23.8302 | -107.078
23.8338 | -107.0785
23.8374 | -107.0796
23.841 | -107.0804
23.8446 | -107.0791
23.8482 | -107.0807
23.8519 | -107.0832
23.8555 | -107.0845
23.8591 | -107.0845
23.8627 | -107.0842
23.8663 | -107.0859
23.8699 | -107.0838
23.8735 | -107.0833
23.8771 | -107.0844
23.8808 | -107.086
23.8844 | -107.0847
23.888 | -107.0827
23.8916 | -107.0829
23.8952 | -107.0814
23.8988 | -107.0785
23.9024 | -107.078
23.9061 | -107.0755
23.9097 | -107.0747
23.9133 | -107.0733
23.9169 | -107.0709
23.9205 | -107.0678
23.9241 | -107.0654
23.9277 | -107.0613
23.9313 | -107.0576
23.935 | -107.0547
23.9386 | -107.0494
23.9422 | -107.0458
23.9458 | -107.0427
23.9494 | -107.0389
23.953 | -107.0331
23.9566 | -107.0278
23.9603 | -107.0225
23.9639 | -107.0185
23.9675 | -107.0124
23.9711 | -107.0059
23.9747 | -107.0015
23.9783 | -106.9952
23.9819 | -106.9899
23.9855 | -106.9842
23.9892 | -106.977
23.9928 | -106.971
23.9964 | -106.9645
24 | -106.9589
24.0036 | -106.9543
24.0072 | -106.9511
24.0108 | -106.9461
24.0145 | -106.9414
24.0181 | -106.9375
24.0217 | -106.9338
24.0253 | -106.929
24.0289 | -106.9254
24.0325 | -106.924
24.0361 | -106.921
24.0397 | -106.9164
24.0434 | -106.9136
24.047 | -106.9098
24.0506 | -106.9067
24.0542 | -106.9048
24.0578 | -106.9043
24.0614 | -106.9046
24.065 | -106.9022
24.0687 | -106.8994
24.0723 | -106.8984
24.0759 | -106.8986
24.0795 | -106.8992
24.0831 | -106.9009
24.0867 | -106.9032
24.0903 | -106.9054
24.0939 | -106.9069
24.0976 | -106.9072
24.1012 | -106.9089
24.1048 | -106.9118
24.1084 | -106.915
24.112 | -106.9186
24.1156 | -106.9231
24.1192 | -106.9253
24.1229 | -106.9294
24.1265 | -106.9342
24.1301 | -106.9382
24.1337 | -106.9422
24.1373 | -106.9476
24.1409 | -106.9525
24.1445 | -106.9572
24.1481 | -106.9593
24.1518 | -106.9598
24.1554 | -106.9622
24.159 | -106.9682
24.1626 | -106.973
24.1662 | -106.9794
24.1698 | -106.9855
24.1734 | -106.9891
24.1771 | -106.9918
24.1807 | -106.9943
24.1843 | -106.9973
24.1879 | -106.9985
24.1915 | -107.0006
24.1951 | -107.001
24.1987 | -107.0025
24.2023 | -107.0059
24.206 | -107.0081
24.2096 | -107.0123
24.2132 | -107.0135
24.2168 | -107.0148
24.2204 | -107.0167
24.224 | -107.0186
24.2276 | -107.0218
24.2313 | -107.0233
24.2349 | -107.025
24.2385 | -107.0291
24.2421 | -107.0316
24.2457 | -107.0355
24.2493 | -107.0387
24.2529 | -107.0407
24.2565 | -107.0409
24.2602 | -107.0375
24.2638 | -107.0286
24.2674 | -107.023
24.271 | -107.0161
24.2746 | -107.0093
24.2782 | -107.0021
24.2818 | -106.9935
24.2855 | -106.9839
24.2891 | -106.9726
24.2927 | -106.9615
24.2963 | -106.9495
24.2999 | -106.9347
24.3035 | -106.9194
24.3071 | -106.9042
24.3107 | -106.8895
24.3144 | -106.8745
24.318 | -106.8588
24.3216 | -106.8404
24.3252 | -106.821
24.3288 | -106.8
24.3324 | -106.7783
24.336 | -106.7546
24.3397 | -106.7293
24.3433 | -106.7033
24.3469 | -106.6759
24.3505 | -106.647
24.3541 | -106.6162
24.3577 | -106.5844
24.3613 | -106.552
24.365 | -106.5183
24.3686 | -106.4853
24.3722 | -106.4517
24.3758 | -106.4184
24.3794 | -106.3856
24.383 | -106.3536
24.3866 | -106.3202
24.3902 | -106.289
24.3939 | -106.2588
24.3975 | -106.2271
24.4011 | -106.1963
24.4047 | -106.1668
24.4083 | -106.1396
24.4119 | -106.1131
24.4155 | -106.0882
24.4192 | -106.0633
24.4228 | -106.039
24.4264 | -106.0157
24.43 | -105.9957
24.4336 | -105.9783
24.4372 | -105.9636
24.4408 | -105.9514
24.4444 | -105.9403
24.4481 | -105.9302
24.4517 | -105.9228
24.4553 | -105.9178
24.4589 | -105.9153
24.4625 | -105.9161
24.4661 | -105.9163
24.4697 | -105.9171
24.4734 | -105.9216
24.477 | -105.9286
24.4806 | -105.9363
24.4842 | -105.9472
24.4878 | -105.9598
24.4914 | -105.9747
24.495 | -105.9928
24.4986 | -106.0193
24.5023 | -106.0362
24.5059 | -106.0533
24.5095 | -106.0719
24.5131 | -106.0905
24.5167 | -106.1097
24.5203 | -106.1303
24.5239 | -106.1512
24.5276 | -106.1729
24.5312 | -106.1944
24.5348 | -106.2164
24.5384 | -106.2363
24.542 | -106.2576
24.5456 | -106.2774
24.5492 | -106.2961
24.5528 | -106.3148
24.5565 | -106.3325
24.5601 | -106.3513
24.5637 | -106.3692
24.5673 | -106.3864
24.5709 | -106.4039
24.5745 | -106.419
24.5781 | -106.4332
24.5818 | -106.4468
24.5854 | -106.46
24.589 | -106.4736
24.5926 | -106.4836
24.5962 | -106.4939
24.5998 | -106.5042
24.6034 | -106.5154
24.607 | -106.5282
24.6107 | -106.5437
24.6143 | -106.5556
24.6179 | -106.5671
24.6215 | -106.5761
24.6251 | -106.5832
24.6287 | -106.5881
24.6323 | -106.592
24.636 | -106.5951
24.6396 | -106.5981
24.6432 | -106.5996
24.6468 | -106.6018
24.6504 | -106.6007
24.654 | -106.5989
24.6576 | -106.5961
24.6612 | -106.5928
24.6649 | -106.5908
24.6685 | -106.5897
24.6721 | -106.5889
24.6757 | -106.5884
24.6793 | -106.5852
24.6829 | -106.5811
24.6865 | -106.5798
24.6902 | -106.579
24.6938 | -106.5773
24.6974 | -106.5762
24.701 | -106.5779
24.7046 | -106.5795
24.7082 | -106.5808
24.7118 | -106.5835
24.7154 | -106.5866
24.7191 | -106.5887
24.7227 | -106.5907
24.7263 | -106.5941
24.7299 | -106.5959
24.7335 | -106.5966
24.7371 | -106.5984
24.7407 | -106.6019
24.7444 | -106.6083
24.748 | -106.6141
24.7516 | -106.6181
24.7552 | -106.6211
24.7588 | -106.6257
24.7624 | -106.6298
24.766 | -106.6351
24.7696 | -106.6423
24.7733 | -106.6493
24.7769 | -106.6542
24.7805 | -106.659
24.7841 | -106.6644
24.7877 | -106.67
24.7913 | -106.6759
24.7949 | -106.6815
24.7986 | -106.6859
24.8022 | -106.6906
24.8058 | -106.6939
24.8094 | -106.6974
24.813 | -106.7026
24.8166 | -106.7091
24.8202 | -106.7128
24.8238 | -106.716
24.8275 | -106.7201
24.8311 | -106.7235
24.8347 | -106.7266
24.8383 | -106.7295
24.8419 | -106.73
24.8455 | -106.7321
24.8491 | -106.7361
24.8528 | -106.7398
24.8564 | -106.7425
24.86 | -106.7447
24.8636 | -106.7479
24.8672 | -106.7514
24.8708 | -106.7535
24.8744 | -106.7558
24.878 | -106.759
24.8817 | -106.7618
24.8853 | -106.7622
24.8889 | -106.7645
24.8925 | -106.7669
24.8961 | -106.7708
24.8997 | -106.7736
24.9033 | -106.7757
24.907 | -106.7779
24.9106 | -106.7798
24.9142 | -106.7813
24.9178 | -106.7808
24.9214 | -106.7811
24.925 | -106.7827
24.9286 | -106.7842
24.9322 | -106.7846
24.9359 | -106.7835
24.9395 | -106.781
24.9431 | -106.7794
24.9467 | -106.7758
24.9503 | -106.7728
24.9539 | -106.7697
24.9575 | -106.7655
24.9612 | -106.7611
24.9648 | -106.7571
24.9684 | -106.7513
24.972 | -106.7463
24.9756 | -106.7401
24.9792 | -106.7335
24.9828 | -106.7279
24.9865 | -106.7227
24.9901 | -106.7155
24.9937 | -106.7067
24.9973 | -106.6978
25.0009 | -106.6915
25.0045 | -106.6853
25.0081 | -106.6778
25.0117 | -106.6695
25.0154 | -106.661
25.019 | -106.6515
25.0226 | -106.6408
25.0262 | -106.6314
25.0298 | -106.6237
25.0334 | -106.6174
25.037 | -106.6111
25.0406 | -106.6029
25.0443 | -106.5941
25.0479 | -106.5863
25.0515 | -106.5795
25.0551 | -106.5724
25.0587 | -106.5663
25.0623 | -106.5613
25.0659 | -106.5555
25.0696 | -106.5498
25.0732 | -106.544
25.0768 | -106.5391
25.0804 | -106.5343
25.084 | -106.5299
25.0876 | -106.5281
25.0912 | -106.5255
25.0949 | -106.5229
25.0985 | -106.5194
25.1021 | -106.5169
25.1057 | -106.5163
25.1093 | -106.5177
25.1129 | -106.5188
25.1165 | -106.5193
25.1201 | -106.5182
25.1238 | -106.5171
25.1274 | -106.5155
25.131 | -106.515
25.1346 | -106.5155
25.1382 | -106.5156
25.1418 | -106.5163
25.1454 | -106.5178
25.1491 | -106.5203
25.1527 | -106.5235
25.1563 | -106.5267
25.1599 | -106.5296
25.1635 | -106.532
25.1671 | -106.5355
25.1707 | -106.5387
25.1743 | -106.5424
25.178 | -106.5472
25.1816 | -106.5512
25.1852 | -106.5546
25.1888 | -106.5583
25.1924 | -106.5616
25.196 | -106.5677
25.1996 | -106.5725
25.2033 | -106.576
25.2069 | -106.5799
25.2105 | -106.5849
25.2141 | -106.5889
25.2177 | -106.591
25.2213 | -106.5943
25.2249 | -106.5967
25.2285 | -106.5988
25.2322 | -106.6012
25.2358 | -106.606
25.2394 | -106.6104
25.243 | -106.6128
25.2466 | -106.615
25.2502 | -106.617
25.2538 | -106.6197
25.2575 | -106.6222
25.2611 | -106.624
25.2647 | -106.6238
25.2683 | -106.6235
25.2719 | -106.624
25.2755 | -106.6246
25.2791 | -106.6234
25.2827 | -106.6225
25.2864 | -106.6214
25.29 | -106.6203
25.2936 | -106.6198
25.2972 | -106.6196
25.3008 | -106.6174
25.3044 | -106.6137
25.308 | -106.61
25.3117 | -106.6061
25.3153 | -106.6022
25.3189 | -106.5974
25.3225 | -106.5927
25.3261 | -106.5865
25.3297 | -106.5805
25.3333 | -106.5761
25.3369 | -106.57
25.3406 | -106.5634
25.3442 | -106.5586
25.3478 | -106.5545
25.3514 | -106.5487
25.355 | -106.5405
25.3586 | -106.532
25.3622 | -106.5244
25.3659 | -106.5169
25.3695 | -106.5075
25.3731 | -106.4987
25.3767 | -106.4927
25.3803 | -106.4857
25.3839 | -106.4797
25.3875 | -106.472
25.3911 | -106.4644
25.3948 | -106.4561
25.3984 | -106.4481
25.402 | -106.4417
25.4056 | -106.4359
25.4092 | -106.4298
25.4128 | -106.4236
25.4164 | -106.4163
25.4201 | -106.4086
25.4237 | -106.4024
25.4273 | -106.3985
25.4309 | -106.3962
25.4345 | -106.3931
25.4381 | -106.3898
25.4417 | -106.3887
25.4453 | -106.3874
25.449 | -106.3863
25.4526 | -106.3856
25.4562 | -106.3862
25.4598 | -106.3869
25.4634 | -106.3875
25.467 | -106.3884
25.4706 | -106.3889
25.4743 | -106.3915
25.4779 | -106.3942
25.4815 | -106.3952
25.4851 | -106.3989
25.4887 | -106.4029
25.4923 | -106.4062
25.4959 | -106.41
25.4995 | -106.4146
25.5032 | -106.4198
25.5068 | -106.425
25.5104 | -106.4304
25.514 | -106.4361
25.5176 | -106.4429
25.5212 | -106.4512
25.5248 | -106.4564
25.5285 | -106.4612
25.5321 | -106.4684
25.5357 | -106.4752
25.5393 | -106.4824
25.5429 | -106.4881
25.5465 | -106.4958
25.5501 | -106.5044
25.5537 | -106.5132
25.5574 | -106.5199
25.561 | -106.5268
25.5646 | -106.5337
25.5682 | -106.54
25.5718 | -106.5459
25.5754 | -106.5523
25.579 | -106.5586
25.5827 | -106.5649
25.5863 | -106.572
25.5899 | -106.5781
25.5935 | -106.584
25.5971 | -106.5876
25.6007 | -106.5926
25.6043 | -106.5976
25.6079 | -106.5993
25.6116 | -106.6014
25.6152 | -106.6039
25.6188 | -106.6063
25.6224 | -106.6074
25.626 | -106.6097
25.6296 | -106.6113
25.6332 | -106.6146
25.6369 | -106.6146
25.6405 | -106.6154
25.6441 | -106.6143
25.6477 | -106.6128
25.6513 | -106.6111
25.6549 | -106.61
25.6585 | -106.6093
25.6622 | -106.6093
25.6658 | -106.6083
25.6694 | -106.6072
25.673 | -106.606
25.6766 | -106.6027
25.6802 | -106.5988
25.6838 | -106.5944
25.6874 | -106.5918
25.6911 | -106.5904
25.6947 | -106.5864
25.6983 | -106.5827
25.7019 | -106.5788
25.7055 | -106.575
25.7091 | -106.5709
25.7127 | -106.568
25.7163 | -106.5638
25.72 | -106.5596
25.7236 | -106.5556
25.7272 | -106.5522
25.7308 | -106.5493
25.7344 | -106.546
25.738 | -106.5427
25.7416 | -106.5391
25.7453 | -106.5347
25.7489 | -106.5315
25.7525 | -106.5278
25.7561 | -106.5248
25.7597 | -106.5234
25.7633 | -106.5203
25.7669 | -106.5162
25.7706 | -106.5138
25.7742 | -106.5102
25.7778 | -106.5084
25.7814 | -106.5066
25.785 | -106.5042
25.7886 | -106.5006
25.7922 | -106.4978
25.7958 | -106.4957
25.7995 | -106.4944
25.8031 | -106.4936
25.8067 | -106.4914
25.8103 | -106.4889
25.8139 | -106.4853
25.8175 | -106.4826
25.8211 | -106.4824
25.8248 | -106.4821
25.8284 | -106.4816
25.832 | -106.4819
25.8356 | -106.4817
25.8392 | -106.4832
25.8428 | -106.4871
25.8464 | -106.4878
25.85 | -106.4882
25.8537 | -106.4901
25.8573 | -106.4919
25.8609 | -106.4931
25.8645 | -106.4946
25.8681 | -106.4961
25.8717 | -106.4982
25.8753 | -106.5006
25.879 | -106.503
25.8826 | -106.5037
25.8862 | -106.5074
25.8898 | -106.5107
25.8934 | -106.5128
25.897 | -106.516
25.9006 | -106.5186
25.9042 | -106.5218
25.9079 | -106.5254
25.9115 | -106.5284
25.9151 | -106.5296
25.9187 | -106.5332
25.9223 | -106.5373
25.9259 | -106.5421
25.9295 | -106.5475
25.9332 | -106.553
25.9368 | -106.5589
25.9404 | -106.5651
25.944 | -106.5722
25.9476 | -106.578
25.9512 | -106.5839
25.9548 | -106.591
25.9584 | -106.5956
25.9621 | -106.6015
25.9657 | -106.6077
25.9693 | -106.6149
25.9729 | -106.6243
25.9765 | -106.6322
25.9801 | -106.6395
25.9837 | -106.648
25.9874 | -106.6585
25.991 | -106.667
25.9946 | -106.6757
25.9982 | -106.6818
26.0018 | -106.6902
26.0054 | -106.6987
26.009 | -106.7093
26.0126 | -106.7196
26.0163 | -106.7286
26.0199 | -106.7366
26.0235 | -106.7447
26.0271 | -106.7523
26.0307 | -106.7604
26.0343 | -106.7692
26.0379 | -106.7781
26.0416 | -106.7884
26.0452 | -106.7971
26.0488 | -106.8057
26.0524 | -106.8138
26.056 | -106.8206
26.0596 | -106.8273
26.0632 | -106.8302
26.0668 | -106.8349
26.0705 | -106.8412
26.0741 | -106.847
26.0777 | -106.8516
26.0813 | -106.8567
26.0849 | -106.8625
26.0885 | -106.8683
26.0921 | -106.8729
26.0958 | -106.8759
26.0994 | -106.8792
26.103 | -106.8823
26.1066 | -106.8839
26.1102 | -106.885
26.1138 | -106.8851
26.1174 | -106.8848
26.121 | -106.8862
26.1247 | -106.8865
26.1283 | -106.8841
26.1319 | -106.8813
26.1355 | -106.88
26.1391 | -106.8757
26.1427 | -106.8727
26.1463 | -106.87
26.15 | -106.8673
26.1536 | -106.8636
26.1572 | -106.8598
26.1608 | -106.8568
26.1644 | -106.8536
26.168 | -106.8477
26.1716 | -106.8404
26.1752 | -106.835
26.1789 | -106.8313
26.1825 | -106.8281
26.1861 | -106.8237
26.1897 | -106.8174
26.1933 | -106.8115
26.1969 | -106.8051
26.2005 | -106.798
26.2042 | -106.792
26.2078 | -106.786
26.2114 | -106.7801
26.215 | -106.775
26.2186 | -106.7703
26.2222 | -106.7643
26.2258 | -106.7585
26.2294 | -106.7535
26.2331 | -106.7485
26.2367 | -106.744
26.2403 | -106.7389
26.2439 | -106.7337
26.2475 | -106.7291
26.2511 | -106.7256
26.2547 | -106.7225
26.2584 | -106.7211
26.262 | -106.7194
26.2656 | -106.7166
26.2692 | -106.7151
26.2728 | -106.7156
26.2764 | -106.715
26.28 | -106.714
26.2837 | -106.7145
26.2873 | -106.7129
26.2909 | -106.7139
26.2945 | -106.7123
26.2981 | -106.7115
26.3017 | -106.7121
26.3053 | -106.7117
26.3089 | -106.7123
26.3126 | -106.713
26.3162 | -106.7128
26.3198 | -106.7127
26.3234 | -106.7137
26.327 | -106.7169
26.3306 | -106.7192
26.3342 | -106.7213
26.3378 | -106.7227
26.3415 | -106.725
26.3451 | -106.7268
26.3487 | -106.7293
26.3523 | -106.7308
26.3559 | -106.7326
26.3595 | -106.7359
26.3631 | -106.7387
26.3668 | -106.7399
26.3704 | -106.741
26.374 | -106.7414
26.3776 | -106.7425
26.3812 | -106.7447
26.3848 | -106.7466
26.3884 | -106.7469
26.3921 | -106.7487
26.3957 | -106.7506
26.3993 | -106.7526
26.4029 | -106.753
26.4065 | -106.7535
26.4101 | -106.755
26.4137 | -106.7571
26.4173 | -106.7578
26.421 | -106.7567
26.4246 | -106.7564
26.4282 | -106.757
26.4318 | -106.7567
26.4354 | -106.7564
26.439 | -106.7569
26.4426 | -106.7574
26.4463 | -106.7568
26.4499 | -106.7543
26.4535 | -106.751
26.4571 | -106.7512
26.4607 | -106.7507
26.4643 | -106.7501
26.4679 | -106.749
26.4715 | -106.7479
26.4752 | -106.7466
26.4788 | -106.7436
26.4824 | -106.7422
26.486 | -106.7418
26.4896 | -106.7399
26.4932 | -106.7376
26.4968 | -106.735
26.5005 | -106.7325
26.5041 | -106.7291
26.5077 | -106.7235
26.5113 | -106.7192
26.5149 | -106.7158
26.5185 | -106.7122
26.5221 | -106.7075
26.5257 | -106.7036
26.5294 | -106.6979
26.533 | -106.6908
26.5366 | -106.6853
26.5402 | -106.6819
26.5438 | -106.6778
26.5474 | -106.6737
26.551 | -106.6688
26.5547 | -106.6617
26.5583 | -106.6543
26.5619 | -106.649
26.5655 | -106.6443
26.5691 | -106.6411
26.5727 | -106.6377
26.5763 | -106.6335
26.5799 | -106.6294
26.5836 | -106.6245
26.5872 | -106.62
26.5908 | -106.6165
26.5944 | -106.613
26.598 | -106.6101
26.6016 | -106.6069
26.6052 | -106.6027
26.6089 | -106.5985
26.6125 | -106.5957
26.6161 | -106.5933
26.6197 | -106.5902
26.6233 | -106.5875
26.6269 | -106.5848
26.6305 | -106.5813
26.6341 | -106.5764
26.6378 | -106.5731
26.6414 | -106.5704
26.645 | -106.5678
26.6486 | -106.5647
26.6522 | -106.562
26.6558 | -106.5607
26.6594 | -106.5586
26.6631 | -106.5572
26.6667 | -106.555
26.6703 | -106.5532
26.6739 | -106.5508
26.6775 | -106.5467
26.6811 | -106.5435
26.6847 | -106.5419
26.6883 | -106.5412
26.692 | -106.5402
26.6956 | -106.5389
26.6992 | -106.536
26.7028 | -106.5325
26.7064 | -106.5303
26.71 | -106.529
26.7136 | -106.5283
26.7173 | -106.5265
26.7209 | -106.5259
26.7245 | -106.5235
26.7281 | -106.5202
26.7317 | -106.5183
26.7353 | -106.5165
26.7389 | -106.5159
26.7425 | -106.5168
26.7462 | -106.5175
26.7498 | -106.5185
26.7534 | -106.5181
26.757 | -106.5168
26.7606 | -106.5163
26.7642 | -106.5176
26.7678 | -106.5177
26.7715 | -106.5161
26.7751 | -106.5151
26.7787 | -106.5133
26.7823 | -106.5107
26.7859 | -106.509
26.7895 | -106.5088
26.7931 | -106.5091
26.7967 | -106.5107
26.8004 | -106.5114
26.804 | -106.5124
26.8076 | -106.513
26.8112 | -106.5143
26.8148 | -106.5159
26.8184 | -106.5164
26.822 | -106.5167
26.8257 | -106.519
26.8293 | -106.522
26.8329 | -106.5227
26.8365 | -106.5246
26.8401 | -106.5273
26.8437 | -106.5288
26.8473 | -106.5298
26.8509 | -106.5313
26.8546 | -106.5342
26.8582 | -106.5367
26.8618 | -106.5382
26.8654 | -106.5405
26.869 | -106.5433
26.8726 | -106.5481
26.8762 | -106.5516
26.8799 | -106.5545
26.8835 | -106.5586
26.8871 | -106.5631
26.8907 | -106.5662
26.8943 | -106.5689
26.8979 | -106.5728
26.9015 | -106.5771
26.9051 | -106.5816
26.9088 | -106.5847
26.9124 | -106.5871
26.916 | -106.5889
26.9196 | -106.5921
26.9232 | -106.5951
26.9268 | -106.5976
26.9304 | -106.6013
26.9341 | -106.6049
26.9377 | -106.6079
26.9413 | -106.6096
26.9449 | -106.6114
26.9485 | -106.6136
26.9521 | -106.615
26.9557 | -106.6162
26.9594 | -106.6193
26.963 | -106.6218
26.9666 | -106.6215
26.9702 | -106.6212
26.9738 | -106.62
26.9774 | -106.6192
26.981 | -106.621
26.9846 | -106.6231
26.9883 | -106.6251
26.9919 | -106.6256
26.9955 | -106.6264
26.9991 | -106.6273
27.0027 | -106.6276
27.0063 | -106.6289
27.0099 | -106.6288
27.0135 | -106.63
27.0172 | -106.6299
27.0208 | -106.6307
27.0244 | -106.6318
27.028 | -106.6327
27.0316 | -106.6326
27.0352 | -106.6309
27.0388 | -106.6316
27.0425 | -106.6322
27.0461 | -106.6323
27.0497 | -106.6331
27.0533 | -106.634
27.0569 | -106.6365
27.0605 | -106.6386
27.0641 | -106.6384
27.0678 | -106.6383
27.0714 | -106.6385
27.075 | -106.6393
27.0786 | -106.6388
27.0822 | -106.638
27.0858 | -106.6371
27.0894 | -106.6355
27.093 | -106.6353
27.0967 | -106.6353
27.1003 | -106.6349
27.1039 | -106.6357
27.1075 | -106.6356
27.1111 | -106.6365
27.1147 | -106.6356
27.1183 | -106.6344
27.122 | -106.6337
27.1256 | -106.6332
27.1292 | -106.6303
27.1328 | -106.6273
27.1364 | -106.6253
27.14 | -106.6233
27.1436 | -106.6221
27.1472 | -106.6194
27.1509 | -106.616
27.1545 | -106.6168
27.1581 | -106.6162
27.1617 | -106.6162
27.1653 | -106.6168
27.1689 | -106.6149
27.1725 | -106.6134
27.1762 | -106.6133
27.1798 | -106.6123
27.1834 | -106.6111
27.187 | -106.6085
27.1906 | -106.6076
27.1942 | -106.6068
27.1978 | -106.6052
27.2014 | -106.6041
27.2051 | -106.6025
27.2087 | -106.6012
27.2123 | -106.5998
27.2159 | -106.5993
27.2195 | -106.5976
27.2231 | -106.5964
27.2267 | -106.5976
27.2304 | -106.5989
27.234 | -106.5988
27.2376 | -106.5993
27.2412 | -106.5997
27.2448 | -106.5999
27.2484 | -106.601
27.252 | -106.6015
27.2556 | -106.6012
27.2593 | -106.5999
27.2629 | -106.6001
27.2665 | -106.6003
27.2701 | -106.6001
27.2737 | -106.6005
27.2773 | -106.601
27.2809 | -106.6021
27.2846 | -106.6028
27.2882 | -106.604
27.2918 | -106.6057
27.2954 | -106.6072
27.299 | -106.6104
27.3026 | -106.6138
27.3062 | -106.6166
27.3098 | -106.6195
27.3135 | -106.6223
27.3171 | -106.6251
27.3207 | -106.6263
27.3243 | -106.6274
27.3279 | -106.6276
27.3315 | -106.6289
27.3351 | -106.6317
27.3388 | -106.6328
27.3424 | -106.6335
27.346 | -106.6356
27.3496 | -106.6393
27.3532 | -106.6415
27.3568 | -106.6435
27.3604 | -106.6448
27.364 | -106.6428
27.3677 | -106.6417
27.3713 | -106.64
27.3749 | -106.6387
27.3785 | -106.638
27.3821 | -106.6382
27.3857 | -106.6371
27.3893 | -106.6353
27.393 | -106.6335
27.3966 | -106.6296
27.4002 | -106.6266
27.4038 | -106.6225
27.4074 | -106.6187
27.411 | -106.615
27.4146 | -106.612
27.4182 | -106.6109
27.4219 | -106.6085
27.4255 | -106.6064
27.4291 | -106.6014
27.4327 | -106.5961
27.4363 | -106.5923
27.4399 | -106.5867
27.4435 | -106.5803
27.4472 | -106.5749
27.4508 | -106.5686
27.4544 | -106.5643
27.458 | -106.5601
27.4616 | -106.5567
27.4652 | -106.5541
27.4688 | -106.5503
27.4724 | -106.5461
27.4761 | -106.541
27.4797 | -106.5361
27.4833 | -106.5318
27.4869 | -106.5278
27.4905 | -106.5239
27.4941 | -106.5218
27.4977 | -106.5195
27.5014 | -106.5167
27.505 | -106.5145
27.5086 | -106.5121
27.5122 | -106.5079
27.5158 | -106.5062
27.5194 | -106.5036
27.523 | -106.4998
27.5266 | -106.497
27.5303 | -106.4925
27.5339 | -106.4889
27.5375 | -106.4858
27.5411 | -106.4816
27.5447 | -106.4784
27.5483 | -106.4744
27.5519 | -106.471
27.5556 | -106.4684
27.5592 | -106.4649
27.5628 | -106.4599
27.5664 | -106.4556
27.57 | -106.4524
27.5736 | -106.4498
27.5772 | -106.4465
27.5808 | -106.4439
27.5845 | -106.4411
27.5881 | -106.438
27.5917 | -106.4369
27.5953 | -106.4362
27.5989 | -106.4347
27.6025 | -106.4327
27.6061 | -106.4313
27.6098 | -106.4297
27.6134 | -106.4283
27.617 | -106.4279
27.6206 | -106.4292
27.6242 | -106.4303
27.6278 | -106.4332
27.6314 | -106.4329
27.635 | -106.4335
27.6387 | -106.4348
27.6423 | -106.4357
27.6459 | -106.4369
27.6495 | -106.4413
27.6531 | -106.4456
27.6567 | -106.4493
27.6603 | -106.4507
27.664 | -106.4516
27.6676 | -106.4542
27.6712 | -106.4568
27.6748 | -106.4598
27.6784 | -106.4646
27.682 | -106.4711
27.6856 | -106.4761
27.6893 | -106.479
27.6929 | -106.4814
27.6965 | -106.4851
27.7001 | -106.4896
27.7037 | -106.4955
27.7073 | -106.5009
27.7109 | -106.5064
27.7145 | -106.5139
27.7182 | -106.5217
27.7218 | -106.5289
27.7254 | -106.5363
27.729 | -106.5439
27.7326 | -106.5513
27.7362 | -106.5578
27.7398 | -106.5649
27.7435 | -106.5731
27.7471 | -106.5821
27.7507 | -106.5909
27.7543 | -106.6004
27.7579 | -106.6096
27.7615 | -106.6188
27.7651 | -106.6274
27.7687 | -106.6345
27.7724 | -106.6404
27.776 | -106.647
27.7796 | -106.6538
27.7832 | -106.661
27.7868 | -106.6694
27.7904 | -106.6774
27.794 | -106.684
27.7977 | -106.6907
27.8013 | -106.6959
27.8049 | -106.7013
27.8085 | -106.7065
27.8121 | -106.7101
27.8157 | -106.7136
27.8193 | -106.7209
27.8229 | -106.7279
27.8266 | -106.7312
27.8302 | -106.7343
27.8338 | -106.7384
27.8374 | -106.74
27.841 | -106.7408
27.8446 | -106.7409
27.8482 | -106.7414
27.8519 | -106.742
27.8555 | -106.7404
27.8591 | -106.7383
27.8627 | -106.737
27.8663 | -106.7352
27.8699 | -106.7334
27.8735 | -106.7322
27.8771 | -106.7312
27.8808 | -106.7282
27.8844 | -106.7233
27.888 | -106.7202
27.8916 | -106.7164
27.8952 | -106.7107
27.8988 | -106.7057
27.9024 | -106.7024
27.9061 | -106.6983
27.9097 | -106.6946
27.9133 | -106.6901
27.9169 | -106.6853
27.9205 | -106.6804
27.9241 | -106.6749
27.9277 | -106.6694
27.9313 | -106.6643
27.935 | -106.6585
27.9386 | -106.6524
27.9422 | -106.6456
27.9458 | -106.6384
27.9494 | -106.6317
27.953 | -106.6276
27.9566 | -106.6229
27.9603 | -106.6164
27.9639 | -106.6099
27.9675 | -106.6044
27.9711 | -106.5997
27.9747 | -106.5953
27.9783 | -106.591
27.9819 | -106.5866
27.9855 | -106.5831
27.9892 | -106.5797
27.9928 | -106.577
27.9964 | -106.5743
28 | -106.5722
28.0036 | -106.5703