## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 223 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 223
Hood | 02-92B/P | 162
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1015.5 | hPa |  | PRC PRO 01
Humidity | 45.5 | %/Hr
Internal temperature of calibration bench | 23.86 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1018 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.3 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.99 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.01 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3100 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 07/05/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 14/05/2018
Engineer | Anthony HUE | Approval |  | 18/05/2018
CEO | Robert GLEMOT | Approval |  | 18/05/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4215 | 58.3587
0.012209 | -36.0823 | 39.2712
0.018313 | -35.1366 | 28.9158
0.024417 | -34.6907 | 22.2185
0.030521 | -34.4728 | 18.0874
0.036626 | -34.3234 | 15.0427
0.04273 | -34.243 | 13.0209
0.048834 | -34.1811 | 11.2558
0.054938 | -34.1574 | 9.8864
0.061043 | -34.1322 | 8.8716
0.067147 | -34.1133 | 8.0181
0.073251 | -34.0921 | 7.2793
0.079355 | -34.0685 | 6.7436
0.08546 | -34.0545 | 6.3512
0.091564 | -34.0495 | 5.9502
0.097668 | -34.0371 | 5.4936
0.10377 | -34.0172 | 5.042
0.10988 | -34.0066 | 4.629
0.11598 | -34.0068 | 4.2779
0.12209 | -34.0183 | 4.0249
0.12819 | -34.0348 | 3.8465
0.13429 | -34.0485 | 3.6989
0.1404 | -34.0581 | 3.5575
0.1465 | -34.0684 | 3.4058
0.15261 | -34.0775 | 3.2422
0.15871 | -34.0824 | 3.058
0.16482 | -34.0797 | 2.8412
0.17092 | -34.0599 | 2.652
0.17702 | -34.0449 | 2.4753
0.18313 | -34.0297 | 2.3333
0.18923 | -34.0182 | 2.2216
0.19534 | -34.0134 | 2.1473
0.20144 | -34.0143 | 2.0927
0.20754 | -34.0216 | 2.0617
0.21365 | -34.0258 | 2.0197
0.21975 | -34.0313 | 1.9704
0.22586 | -34.0384 | 1.9114
0.23196 | -34.044 | 1.8286
0.23807 | -34.0476 | 1.7312
0.24417 | -34.0445 | 1.6211
0.25027 | -34.043 | 1.5241
0.25638 | -34.0401 | 1.4231
0.26248 | -34.0363 | 1.3379
0.26859 | -34.0334 | 1.2585
0.27469 | -34.0329 | 1.188
0.2808 | -34.0381 | 1.1521
0.2869 | -34.0422 | 1.1065
0.293 | -34.0471 | 1.0697
0.29911 | -34.0528 | 1.0448
0.30521 | -34.058 | 1.0227
0.31132 | -34.0616 | 0.99982
0.31742 | -34.0648 | 0.97322
0.32353 | -34.0684 | 0.94136
0.32963 | -34.0707 | 0.90572
0.33573 | -34.0718 | 0.86101
0.34184 | -34.0731 | 0.81567
0.34794 | -34.0741 | 0.75807
0.35405 | -34.0751 | 0.70524
0.36015 | -34.0758 | 0.65174
0.36626 | -34.0761 | 0.59891
0.37236 | -34.0762 | 0.54764
0.37846 | -34.0761 | 0.50069
0.38457 | -34.0744 | 0.47397
0.39067 | -34.0741 | 0.44008
0.39678 | -34.0721 | 0.4088
0.40288 | -34.0696 | 0.38489
0.40899 | -34.0664 | 0.36022
0.41509 | -34.0634 | 0.33504
0.42119 | -34.0588 | 0.3065
0.4273 | -34.0556 | 0.2793
0.4334 | -34.0528 | 0.25197
0.43951 | -34.0498 | 0.2224
0.44561 | -34.0476 | 0.19147
0.45172 | -34.0475 | 0.15927
0.45782 | -34.0466 | 0.12972
0.46392 | -34.0467 | 0.10292
0.47003 | -34.0475 | 0.074925
0.47613 | -34.049 | 0.053116
0.48224 | -34.0511 | 0.034551
0.48834 | -34.0556 | 0.027036
0.49445 | -34.0587 | 0.016177
0.50055 | -34.0627 | 0.003592
0.50665 | -34.0662 | -0.0011305
0.51276 | -34.0695 | -0.0071428
0.51886 | -34.0728 | -0.01009
0.52497 | -34.0742 | -0.014528
0.53107 | -34.0764 | -0.015926
0.53717 | -34.0784 | -0.019714
0.54328 | -34.08 | -0.024628
0.54938 | -34.0809 | -0.032612
0.55549 | -34.081 | -0.043192
0.56159 | -34.0799 | -0.067953
0.5677 | -34.0796 | -0.088456
0.5738 | -34.0781 | -0.10897
0.5799 | -34.0776 | -0.12989
0.58601 | -34.077 | -0.16187
0.59211 | -34.0761 | -0.20412
0.59822 | -34.0753 | -0.23649
0.60432 | -34.0747 | -0.26992
0.61043 | -34.0743 | -0.30859
0.61653 | -34.0742 | -0.3439
0.62263 | -34.0735 | -0.38181
0.62874 | -34.0744 | -0.41421
0.63484 | -34.0747 | -0.44993
0.64095 | -34.0757 | -0.48403
0.64705 | -34.0768 | -0.51743
0.65316 | -34.078 | -0.54633
0.65926 | -34.0794 | -0.5733
0.66536 | -34.0809 | -0.59308
0.67147 | -34.0822 | -0.61725
0.67757 | -34.0837 | -0.64126
0.68368 | -34.0851 | -0.66055
0.68978 | -34.0868 | -0.6802
0.69589 | -34.0883 | -0.69727
0.70199 | -34.0893 | -0.71175
0.70809 | -34.0905 | -0.7301
0.7142 | -34.0916 | -0.74537
0.7203 | -34.0928 | -0.76268
0.72641 | -34.0934 | -0.77625
0.73251 | -34.0946 | -0.79313
0.73862 | -34.0948 | -0.81003
0.74472 | -34.0961 | -0.82492
0.75082 | -34.0965 | -0.83839
0.75693 | -34.0966 | -0.85453
0.76303 | -34.0963 | -0.86912
0.76914 | -34.0947 | -0.88955
0.77524 | -34.094 | -0.90974
0.78135 | -34.0931 | -0.92748
0.78745 | -34.0922 | -0.94685
0.79355 | -34.0915 | -0.96605
0.79966 | -34.0904 | -0.98681
0.80576 | -34.0888 | -1.0113
0.81187 | -34.0873 | -1.0326
0.81797 | -34.0856 | -1.0553
0.82408 | -34.0843 | -1.0755
0.83018 | -34.0828 | -1.0995
0.83628 | -34.0815 | -1.1259
0.84239 | -34.0801 | -1.1467
0.84849 | -34.0785 | -1.166
0.8546 | -34.0773 | -1.1875
0.8607 | -34.076 | -1.2091
0.86681 | -34.0751 | -1.2323
0.87291 | -34.0736 | -1.253
0.87901 | -34.0729 | -1.2713
0.88512 | -34.0721 | -1.2902
0.89122 | -34.0714 | -1.3091
0.89733 | -34.071 | -1.3277
0.90343 | -34.0705 | -1.3455
0.90953 | -34.0709 | -1.3609
0.91564 | -34.0708 | -1.3757
0.92174 | -34.0705 | -1.3898
0.92785 | -34.0703 | -1.4037
0.93395 | -34.0706 | -1.4168
0.94006 | -34.0708 | -1.4331
0.94616 | -34.0714 | -1.4413
0.95226 | -34.0718 | -1.4555
0.95837 | -34.0722 | -1.4681
0.96447 | -34.0727 | -1.4806
0.97058 | -34.0734 | -1.492
0.97668 | -34.0739 | -1.5014
0.98279 | -34.0747 | -1.5067
0.98889 | -34.0756 | -1.5166
0.99499 | -34.0761 | -1.5262
1.0011 | -34.077 | -1.5359
1.0072 | -34.0778 | -1.543
1.0133 | -34.0786 | -1.5515
1.0194 | -34.0795 | -1.558
1.0255 | -34.0801 | -1.5652
1.0316 | -34.0806 | -1.5721
1.0377 | -34.0811 | -1.5781
1.0438 | -34.0816 | -1.5841
1.0499 | -34.0815 | -1.5879
1.056 | -34.0817 | -1.5919
1.0621 | -34.0818 | -1.5967
1.0682 | -34.0823 | -1.6032
1.0743 | -34.0825 | -1.6084
1.0805 | -34.083 | -1.6139
1.0866 | -34.083 | -1.6222
1.0927 | -34.0832 | -1.6281
1.0988 | -34.0836 | -1.6367
1.1049 | -34.0836 | -1.6431
1.111 | -34.0835 | -1.6514
1.1171 | -34.0835 | -1.6589
1.1232 | -34.0832 | -1.6701
1.1293 | -34.0829 | -1.6809
1.1354 | -34.0827 | -1.6907
1.1415 | -34.0824 | -1.7014
1.1476 | -34.0822 | -1.7122
1.1537 | -34.0819 | -1.7232
1.1598 | -34.082 | -1.7389
1.1659 | -34.0821 | -1.7506
1.172 | -34.0821 | -1.7626
1.1781 | -34.082 | -1.7754
1.1842 | -34.082 | -1.7894
1.1903 | -34.0826 | -1.8072
1.1964 | -34.0829 | -1.8211
1.2025 | -34.083 | -1.835
1.2086 | -34.0834 | -1.8513
1.2147 | -34.0839 | -1.8668
1.2209 | -34.0845 | -1.8826
1.227 | -34.0857 | -1.9004
1.2331 | -34.0864 | -1.9159
1.2392 | -34.0872 | -1.9332
1.2453 | -34.0881 | -1.9476
1.2514 | -34.0892 | -1.9638
1.2575 | -34.0903 | -1.9804
1.2636 | -34.0917 | -1.9971
1.2697 | -34.0929 | -2.0137
1.2758 | -34.0939 | -2.0301
1.2819 | -34.095 | -2.0476
1.288 | -34.0961 | -2.064
1.2941 | -34.0972 | -2.0805
1.3002 | -34.0986 | -2.0951
1.3063 | -34.0998 | -2.1101
1.3124 | -34.101 | -2.1251
1.3185 | -34.1023 | -2.1408
1.3246 | -34.1034 | -2.1565
1.3307 | -34.1047 | -2.1715
1.3368 | -34.1057 | -2.1842
1.3429 | -34.1067 | -2.198
1.349 | -34.1078 | -2.211
1.3551 | -34.1088 | -2.2263
1.3613 | -34.1099 | -2.2401
1.3674 | -34.1106 | -2.2517
1.3735 | -34.1115 | -2.2635
1.3796 | -34.1124 | -2.2757
1.3857 | -34.1133 | -2.2871
1.3918 | -34.1139 | -2.2999
1.3979 | -34.1146 | -2.3128
1.404 | -34.1149 | -2.3239
1.4101 | -34.1155 | -2.3356
1.4162 | -34.1161 | -2.3466
1.4223 | -34.1165 | -2.3581
1.4284 | -34.1169 | -2.3695
1.4345 | -34.1175 | -2.3808
1.4406 | -34.1179 | -2.3899
1.4467 | -34.1182 | -2.4
1.4528 | -34.1185 | -2.4103
1.4589 | -34.1189 | -2.4212
1.465 | -34.1191 | -2.4307
1.4711 | -34.1196 | -2.4406
1.4772 | -34.1197 | -2.4508
1.4833 | -34.12 | -2.4617
1.4894 | -34.1203 | -2.4728
1.4955 | -34.1205 | -2.4844
1.5016 | -34.1207 | -2.4957
1.5078 | -34.1209 | -2.5075
1.5139 | -34.1213 | -2.5186
1.52 | -34.1215 | -2.5309
1.5261 | -34.1218 | -2.5416
1.5322 | -34.1218 | -2.5513
1.5383 | -34.1218 | -2.5626
1.5444 | -34.122 | -2.5747
1.5505 | -34.1221 | -2.5857
1.5566 | -34.1222 | -2.5969
1.5627 | -34.1224 | -2.6075
1.5688 | -34.1226 | -2.6181
1.5749 | -34.1227 | -2.6293
1.581 | -34.1229 | -2.6438
1.5871 | -34.1232 | -2.6567
1.5932 | -34.1235 | -2.6681
1.5993 | -34.1235 | -2.6815
1.6054 | -34.1237 | -2.6944
1.6115 | -34.1237 | -2.708
1.6176 | -34.124 | -2.7232
1.6237 | -34.1243 | -2.735
1.6298 | -34.1245 | -2.7492
1.6359 | -34.1246 | -2.7623
1.642 | -34.1248 | -2.7763
1.6482 | -34.1252 | -2.7915
1.6543 | -34.1256 | -2.8054
1.6604 | -34.1258 | -2.8196
1.6665 | -34.1261 | -2.8335
1.6726 | -34.1264 | -2.8486
1.6787 | -34.1267 | -2.8642
1.6848 | -34.127 | -2.8794
1.6909 | -34.1273 | -2.8941
1.697 | -34.1275 | -2.9095
1.7031 | -34.1278 | -2.9239
1.7092 | -34.1282 | -2.9381
1.7153 | -34.1284 | -2.9539
1.7214 | -34.1286 | -2.9685
1.7275 | -34.1288 | -2.983
1.7336 | -34.129 | -2.9975
1.7397 | -34.1291 | -3.0117
1.7458 | -34.1295 | -3.0275
1.7519 | -34.1296 | -3.0429
1.758 | -34.1297 | -3.0571
1.7641 | -34.1299 | -3.0716
1.7702 | -34.13 | -3.0855
1.7763 | -34.1304 | -3.0992
1.7824 | -34.1306 | -3.1143
1.7885 | -34.1307 | -3.127
1.7947 | -34.1309 | -3.1412
1.8008 | -34.1311 | -3.1545
1.8069 | -34.1313 | -3.168
1.813 | -34.1315 | -3.1812
1.8191 | -34.1315 | -3.1939
1.8252 | -34.1315 | -3.206
1.8313 | -34.1316 | -3.2194
1.8374 | -34.1317 | -3.2324
1.8435 | -34.1319 | -3.244
1.8496 | -34.1318 | -3.2563
1.8557 | -34.1318 | -3.2693
1.8618 | -34.1316 | -3.28
1.8679 | -34.1317 | -3.2923
1.874 | -34.1319 | -3.3047
1.8801 | -34.1318 | -3.3164
1.8862 | -34.1318 | -3.3284
1.8923 | -34.1319 | -3.3396
1.8984 | -34.1316 | -3.3494
1.9045 | -34.1317 | -3.3603
1.9106 | -34.1316 | -3.3712
1.9167 | -34.1315 | -3.3814
1.9228 | -34.1315 | -3.3917
1.9289 | -34.1313 | -3.4
1.9351 | -34.1314 | -3.4099
1.9412 | -34.1314 | -3.4204
1.9473 | -34.1313 | -3.4314
1.9534 | -34.1312 | -3.4425
1.9595 | -34.1309 | -3.4535
1.9656 | -34.1303 | -3.463
1.9717 | -34.1301 | -3.473
1.9778 | -34.1299 | -3.4845
1.9839 | -34.1298 | -3.4958
1.99 | -34.1298 | -3.5061
1.9961 | -34.1297 | -3.5155
2.0022 | -34.1295 | -3.5254
2.0083 | -34.1293 | -3.5356
2.0144 | -34.1292 | -3.5454
2.0205 | -34.1291 | -3.5566
2.0266 | -34.1289 | -3.5681
2.0327 | -34.1289 | -3.579
2.0388 | -34.1288 | -3.5924
2.0449 | -34.1289 | -3.6027
2.051 | -34.1288 | -3.6127
2.0571 | -34.1284 | -3.624
2.0632 | -34.1283 | -3.6343
2.0693 | -34.1283 | -3.6443
2.0754 | -34.1284 | -3.656
2.0816 | -34.1284 | -3.6668
2.0877 | -34.1283 | -3.6776
2.0938 | -34.1283 | -3.6888
2.0999 | -34.1283 | -3.6999
2.106 | -34.1285 | -3.7114
2.1121 | -34.1286 | -3.7227
2.1182 | -34.1286 | -3.7334
2.1243 | -34.1287 | -3.7447
2.1304 | -34.1288 | -3.7552
2.1365 | -34.129 | -3.7661
2.1426 | -34.1291 | -3.7784
2.1487 | -34.1294 | -3.7901
2.1548 | -34.1295 | -3.8023
2.1609 | -34.1298 | -3.814
2.167 | -34.13 | -3.8266
2.1731 | -34.1301 | -3.8374
2.1792 | -34.1308 | -3.849
2.1853 | -34.1311 | -3.8609
2.1914 | -34.1315 | -3.873
2.1975 | -34.1318 | -3.8847
2.2036 | -34.1322 | -3.8968
2.2097 | -34.1326 | -3.9085
2.2158 | -34.1333 | -3.9195
2.222 | -34.1338 | -3.9308
2.2281 | -34.134 | -3.9414
2.2342 | -34.1344 | -3.9543
2.2403 | -34.1349 | -3.9653
2.2464 | -34.1356 | -3.9769
2.2525 | -34.1359 | -3.9875
2.2586 | -34.1367 | -3.9982
2.2647 | -34.1373 | -4.0085
2.2708 | -34.1378 | -4.0201
2.2769 | -34.1385 | -4.0324
2.283 | -34.139 | -4.0419
2.2891 | -34.1395 | -4.0526
2.2952 | -34.1402 | -4.0634
2.3013 | -34.1408 | -4.074
2.3074 | -34.1414 | -4.0847
2.3135 | -34.1419 | -4.0958
2.3196 | -34.1425 | -4.1058
2.3257 | -34.1429 | -4.1172
2.3318 | -34.1433 | -4.1265
2.3379 | -34.1438 | -4.1385
2.344 | -34.1443 | -4.1492
2.3501 | -34.1447 | -4.1606
2.3562 | -34.1451 | -4.1699
2.3623 | -34.1455 | -4.1798
2.3685 | -34.1459 | -4.1898
2.3746 | -34.1462 | -4.2002
2.3807 | -34.1467 | -4.2101
2.3868 | -34.1469 | -4.2197
2.3929 | -34.1473 | -4.2297
2.399 | -34.1478 | -4.2388
2.4051 | -34.1482 | -4.2491
2.4112 | -34.1485 | -4.2588
2.4173 | -34.1489 | -4.2685
2.4234 | -34.1491 | -4.2775
2.4295 | -34.1494 | -4.2864
2.4356 | -34.1496 | -4.2962
2.4417 | -34.1497 | -4.3063
2.4478 | -34.1499 | -4.3162
2.4539 | -34.1503 | -4.3253
2.46 | -34.1501 | -4.3345
2.4661 | -34.1502 | -4.3438
2.4722 | -34.1505 | -4.3538
2.4783 | -34.1507 | -4.3631
2.4844 | -34.1509 | -4.3718
2.4905 | -34.151 | -4.3813
2.4966 | -34.1507 | -4.3906
2.5027 | -34.1508 | -4.3995
2.5089 | -34.1508 | -4.4083
2.515 | -34.1508 | -4.418
2.5211 | -34.1509 | -4.4267
2.5272 | -34.1507 | -4.4359
2.5333 | -34.1506 | -4.4456
2.5394 | -34.1506 | -4.455
2.5455 | -34.1505 | -4.464
2.5516 | -34.1505 | -4.4737
2.5577 | -34.1505 | -4.4831
2.5638 | -34.1503 | -4.4923
2.5699 | -34.1503 | -4.5021
2.576 | -34.1503 | -4.5117
2.5821 | -34.1502 | -4.5214
2.5882 | -34.1501 | -4.5308
2.5943 | -34.1501 | -4.5409
2.6004 | -34.1498 | -4.5508
2.6065 | -34.1498 | -4.5601
2.6126 | -34.1498 | -4.57
2.6187 | -34.1497 | -4.5794
2.6248 | -34.1497 | -4.5887
2.6309 | -34.1495 | -4.5993
2.637 | -34.1495 | -4.6094
2.6431 | -34.1495 | -4.6196
2.6492 | -34.1495 | -4.6291
2.6554 | -34.1495 | -4.6395
2.6615 | -34.1494 | -4.6493
2.6676 | -34.1492 | -4.6591
2.6737 | -34.1492 | -4.6694
2.6798 | -34.1491 | -4.679
2.6859 | -34.1491 | -4.689
2.692 | -34.149 | -4.6988
2.6981 | -34.149 | -4.7097
2.7042 | -34.1491 | -4.7215
2.7103 | -34.149 | -4.732
2.7164 | -34.1491 | -4.7425
2.7225 | -34.1491 | -4.7532
2.7286 | -34.1491 | -4.7642
2.7347 | -34.1491 | -4.7748
2.7408 | -34.1492 | -4.785
2.7469 | -34.1492 | -4.7959
2.753 | -34.1493 | -4.8069
2.7591 | -34.1493 | -4.8183
2.7652 | -34.1492 | -4.8289
2.7713 | -34.1493 | -4.839
2.7774 | -34.1496 | -4.8503
2.7835 | -34.1498 | -4.8609
2.7896 | -34.1498 | -4.8707
2.7958 | -34.1501 | -4.8811
2.8019 | -34.1505 | -4.8915
2.808 | -34.1507 | -4.9034
2.8141 | -34.1509 | -4.9159
2.8202 | -34.1511 | -4.928
2.8263 | -34.1514 | -4.9388
2.8324 | -34.1515 | -4.9505
2.8385 | -34.1516 | -4.9625
2.8446 | -34.1519 | -4.9742
2.8507 | -34.152 | -4.9863
2.8568 | -34.1523 | -4.9979
2.8629 | -34.1526 | -5.0096
2.869 | -34.1529 | -5.0205
2.8751 | -34.1531 | -5.0327
2.8812 | -34.1535 | -5.0452
2.8873 | -34.1537 | -5.0564
2.8934 | -34.1539 | -5.068
2.8995 | -34.154 | -5.0795
2.9056 | -34.1543 | -5.0899
2.9117 | -34.1545 | -5.1011
2.9178 | -34.155 | -5.1138
2.9239 | -34.1551 | -5.1261
2.93 | -34.1553 | -5.1386
2.9361 | -34.1558 | -5.1501
2.9423 | -34.1562 | -5.1623
2.9484 | -34.1564 | -5.1754
2.9545 | -34.1568 | -5.1883
2.9606 | -34.157 | -5.2008
2.9667 | -34.1574 | -5.2122
2.9728 | -34.1576 | -5.2238
2.9789 | -34.1577 | -5.236
2.985 | -34.1578 | -5.249
2.9911 | -34.1581 | -5.2609
2.9972 | -34.1583 | -5.2731
3.0033 | -34.1586 | -5.2852
3.0094 | -34.1589 | -5.2978
3.0155 | -34.1592 | -5.3102
3.0216 | -34.1595 | -5.3229
3.0277 | -34.1598 | -5.3354
3.0338 | -34.1601 | -5.3481
3.0399 | -34.1604 | -5.361
3.046 | -34.1606 | -5.373
3.0521 | -34.161 | -5.3851
3.0582 | -34.1612 | -5.3973
3.0643 | -34.1616 | -5.4095
3.0704 | -34.1617 | -5.4214
3.0765 | -34.162 | -5.4333
3.0827 | -34.1622 | -5.4459
3.0888 | -34.1623 | -5.4581
3.0949 | -34.1624 | -5.4703
3.101 | -34.1627 | -5.4835
3.1071 | -34.1628 | -5.4963
3.1132 | -34.1631 | -5.5089
3.1193 | -34.1633 | -5.521
3.1254 | -34.1635 | -5.5323
3.1315 | -34.1637 | -5.5451
3.1376 | -34.1639 | -5.5572
3.1437 | -34.164 | -5.5696
3.1498 | -34.1642 | -5.582
3.1559 | -34.1644 | -5.5944
3.162 | -34.1644 | -5.6067
3.1681 | -34.1644 | -5.6194
3.1742 | -34.1647 | -5.632
3.1803 | -34.1649 | -5.644
3.1864 | -34.1651 | -5.6556
3.1925 | -34.1653 | -5.6683
3.1986 | -34.1653 | -5.6795
3.2047 | -34.1654 | -5.6918
3.2108 | -34.1656 | -5.7031
3.2169 | -34.1659 | -5.7156
3.223 | -34.1658 | -5.7279
3.2292 | -34.1658 | -5.7404
3.2353 | -34.1657 | -5.75
3.2414 | -34.1658 | -5.7621
3.2475 | -34.1658 | -5.7723
3.2536 | -34.1659 | -5.7852
3.2597 | -34.1661 | -5.7962
3.2658 | -34.1663 | -5.8068
3.2719 | -34.1664 | -5.8175
3.278 | -34.1664 | -5.8285
3.2841 | -34.1665 | -5.8401
3.2902 | -34.1666 | -5.8517
3.2963 | -34.1667 | -5.8625
3.3024 | -34.1667 | -5.8724
3.3085 | -34.1669 | -5.8838
3.3146 | -34.1669 | -5.8942
3.3207 | -34.167 | -5.9054
3.3268 | -34.1672 | -5.9167
3.3329 | -34.1673 | -5.9267
3.339 | -34.1675 | -5.9363
3.3451 | -34.1675 | -5.9472
3.3512 | -34.1677 | -5.9568
3.3573 | -34.1678 | -5.9676
3.3634 | -34.168 | -5.9786
3.3696 | -34.1679 | -5.9897
3.3757 | -34.1681 | -5.9996
3.3818 | -34.1683 | -6.0111
3.3879 | -34.1683 | -6.0211
3.394 | -34.1684 | -6.0308
3.4001 | -34.1684 | -6.0416
3.4062 | -34.1686 | -6.0516
3.4123 | -34.1688 | -6.0617
3.4184 | -34.1688 | -6.072
3.4245 | -34.1689 | -6.0825
3.4306 | -34.169 | -6.0924
3.4367 | -34.1692 | -6.1025
3.4428 | -34.1695 | -6.1121
3.4489 | -34.1696 | -6.123
3.455 | -34.1696 | -6.1333
3.4611 | -34.1697 | -6.1437
3.4672 | -34.1697 | -6.1546
3.4733 | -34.1699 | -6.1643
3.4794 | -34.1699 | -6.1746
3.4855 | -34.1702 | -6.1847
3.4916 | -34.1704 | -6.1949
3.4977 | -34.1704 | -6.2051
3.5038 | -34.1707 | -6.2149
3.5099 | -34.1708 | -6.2252
3.5161 | -34.171 | -6.2351
3.5222 | -34.1713 | -6.2457
3.5283 | -34.1714 | -6.2557
3.5344 | -34.1715 | -6.2664
3.5405 | -34.1717 | -6.2765
3.5466 | -34.1719 | -6.2859
3.5527 | -34.1721 | -6.2964
3.5588 | -34.1724 | -6.3064
3.5649 | -34.1726 | -6.3166
3.571 | -34.1728 | -6.327
3.5771 | -34.173 | -6.3372
3.5832 | -34.1732 | -6.3479
3.5893 | -34.1734 | -6.3578
3.5954 | -34.1736 | -6.3681
3.6015 | -34.1738 | -6.3782
3.6076 | -34.174 | -6.388
3.6137 | -34.1742 | -6.3987
3.6198 | -34.1745 | -6.4081
3.6259 | -34.1747 | -6.4185
3.632 | -34.1749 | -6.4291
3.6381 | -34.1752 | -6.4388
3.6442 | -34.1755 | -6.4494
3.6503 | -34.1757 | -6.4591
3.6565 | -34.176 | -6.4705
3.6626 | -34.1762 | -6.4815
3.6687 | -34.1765 | -6.4919
3.6748 | -34.1767 | -6.5021
3.6809 | -34.177 | -6.5125
3.687 | -34.1773 | -6.5232
3.6931 | -34.1776 | -6.5336
3.6992 | -34.1779 | -6.5442
3.7053 | -34.1781 | -6.5544
3.7114 | -34.1784 | -6.5656
3.7175 | -34.1787 | -6.5753
3.7236 | -34.1789 | -6.5847
3.7297 | -34.179 | -6.5952
3.7358 | -34.1793 | -6.6065
3.7419 | -34.1795 | -6.6178
3.748 | -34.1796 | -6.6286
3.7541 | -34.1798 | -6.6395
3.7602 | -34.1801 | -6.6503
3.7663 | -34.1803 | -6.6607
3.7724 | -34.1806 | -6.6715
3.7785 | -34.1808 | -6.682
3.7846 | -34.1811 | -6.6928
3.7907 | -34.1814 | -6.7032
3.7969 | -34.1816 | -6.7145
3.803 | -34.1819 | -6.725
3.8091 | -34.1822 | -6.7354
3.8152 | -34.1824 | -6.7465
3.8213 | -34.1826 | -6.7577
3.8274 | -34.1829 | -6.7684
3.8335 | -34.1831 | -6.7796
3.8396 | -34.1835 | -6.7909
3.8457 | -34.1837 | -6.8008
3.8518 | -34.1841 | -6.8116
3.8579 | -34.1843 | -6.822
3.864 | -34.1846 | -6.8324
3.8701 | -34.1849 | -6.8426
3.8762 | -34.1852 | -6.8539
3.8823 | -34.1854 | -6.8642
3.8884 | -34.1857 | -6.8742
3.8945 | -34.186 | -6.8849
3.9006 | -34.1864 | -6.8958
3.9067 | -34.1867 | -6.9068
3.9128 | -34.187 | -6.9178
3.9189 | -34.1873 | -6.9284
3.925 | -34.1875 | -6.9389
3.9311 | -34.1879 | -6.9493
3.9372 | -34.1882 | -6.9602
3.9434 | -34.1885 | -6.9704
3.9495 | -34.1887 | -6.9816
3.9556 | -34.189 | -6.9927
3.9617 | -34.1892 | -7.0039
3.9678 | -34.1894 | -7.0147
3.9739 | -34.1898 | -7.0257
3.98 | -34.1902 | -7.0373
3.9861 | -34.1904 | -7.0474
3.9922 | -34.1906 | -7.0589
3.9983 | -34.1909 | -7.0692
4.0044 | -34.1911 | -7.0793
4.0105 | -34.1914 | -7.0899
4.0166 | -34.1918 | -7.1005
4.0227 | -34.1922 | -7.1107
4.0288 | -34.1924 | -7.1213
4.0349 | -34.1927 | -7.1313
4.041 | -34.193 | -7.1421
4.0471 | -34.1933 | -7.1521
4.0532 | -34.1936 | -7.1635
4.0593 | -34.1938 | -7.1746
4.0654 | -34.1941 | -7.1855
4.0715 | -34.1944 | -7.196
4.0776 | -34.1946 | -7.2061
4.0838 | -34.1949 | -7.2166
4.0899 | -34.1952 | -7.2281
4.096 | -34.1954 | -7.2383
4.1021 | -34.1957 | -7.2493
4.1082 | -34.1961 | -7.2597
4.1143 | -34.1964 | -7.27
4.1204 | -34.1967 | -7.281
4.1265 | -34.1971 | -7.2916
4.1326 | -34.1974 | -7.3021
4.1387 | -34.1977 | -7.3133
4.1448 | -34.198 | -7.3241
4.1509 | -34.1983 | -7.3347
4.157 | -34.1986 | -7.3452
4.1631 | -34.1988 | -7.356
4.1692 | -34.1991 | -7.3668
4.1753 | -34.1994 | -7.3776
4.1814 | -34.1998 | -7.3879
4.1875 | -34.2 | -7.3988
4.1936 | -34.2003 | -7.4099
4.1997 | -34.2006 | -7.4206
4.2058 | -34.2008 | -7.4317
4.2119 | -34.2011 | -7.4416
4.218 | -34.2014 | -7.4521
4.2241 | -34.2018 | -7.4619
4.2303 | -34.2021 | -7.4733
4.2364 | -34.2025 | -7.4848
4.2425 | -34.2028 | -7.4953
4.2486 | -34.2031 | -7.5063
4.2547 | -34.2034 | -7.516
4.2608 | -34.2037 | -7.5258
4.2669 | -34.204 | -7.5367
4.273 | -34.2043 | -7.5474
4.2791 | -34.2045 | -7.5582
4.2852 | -34.2048 | -7.5683
4.2913 | -34.205 | -7.5781
4.2974 | -34.2053 | -7.5887
4.3035 | -34.2055 | -7.5991
4.3096 | -34.2057 | -7.6094
4.3157 | -34.206 | -7.6197
4.3218 | -34.2063 | -7.63
4.3279 | -34.2066 | -7.6404
4.334 | -34.2068 | -7.6506
4.3401 | -34.2071 | -7.6612
4.3462 | -34.2073 | -7.6718
4.3523 | -34.2077 | -7.6822
4.3584 | -34.2079 | -7.6914
4.3645 | -34.2081 | -7.7014
4.3707 | -34.2084 | -7.7113
4.3768 | -34.2086 | -7.7219
4.3829 | -34.2088 | -7.7319
4.389 | -34.2091 | -7.7419
4.3951 | -34.2094 | -7.7519
4.4012 | -34.2097 | -7.7615
4.4073 | -34.21 | -7.7722
4.4134 | -34.2102 | -7.7826
4.4195 | -34.2105 | -7.7925
4.4256 | -34.2107 | -7.8025
4.4317 | -34.211 | -7.8123
4.4378 | -34.2112 | -7.8227
4.4439 | -34.2114 | -7.8329
4.45 | -34.2117 | -7.8433
4.4561 | -34.2119 | -7.8535
4.4622 | -34.2121 | -7.8634
4.4683 | -34.2123 | -7.8729
4.4744 | -34.2126 | -7.883
4.4805 | -34.2129 | -7.8929
4.4866 | -34.2131 | -7.9025
4.4927 | -34.2133 | -7.9128
4.4988 | -34.2136 | -7.9226
4.5049 | -34.2138 | -7.9326
4.511 | -34.214 | -7.9423
4.5172 | -34.2144 | -7.9525
4.5233 | -34.2147 | -7.9629
4.5294 | -34.2149 | -7.9728
4.5355 | -34.2152 | -7.9825
4.5416 | -34.2155 | -7.992
4.5477 | -34.2157 | -8.0016
4.5538 | -34.2159 | -8.0113
4.5599 | -34.2162 | -8.0212
4.566 | -34.2164 | -8.0312
4.5721 | -34.2167 | -8.0407
4.5782 | -34.217 | -8.0505
4.5843 | -34.2172 | -8.0604
4.5904 | -34.2175 | -8.0698
4.5965 | -34.2178 | -8.0792
4.6026 | -34.2181 | -8.0895
4.6087 | -34.2183 | -8.099
4.6148 | -34.2186 | -8.1096
4.6209 | -34.219 | -8.1185
4.627 | -34.2193 | -8.1285
4.6331 | -34.2195 | -8.138
4.6392 | -34.2198 | -8.1474
4.6453 | -34.22 | -8.1571
4.6514 | -34.2203 | -8.1668
4.6576 | -34.2204 | -8.1771
4.6637 | -34.2207 | -8.1856
4.6698 | -34.221 | -8.1948
4.6759 | -34.2211 | -8.2043
4.682 | -34.2214 | -8.2138
4.6881 | -34.2217 | -8.223
4.6942 | -34.2219 | -8.2326
4.7003 | -34.2222 | -8.2423
4.7064 | -34.2224 | -8.2521
4.7125 | -34.2227 | -8.2613
4.7186 | -34.2229 | -8.2703
4.7247 | -34.2233 | -8.28
4.7308 | -34.2236 | -8.2896
4.7369 | -34.2238 | -8.2994
4.743 | -34.224 | -8.3094
4.7491 | -34.2242 | -8.3188
4.7552 | -34.2244 | -8.3281
4.7613 | -34.2247 | -8.3378
4.7674 | -34.225 | -8.3468
4.7735 | -34.2252 | -8.3566
4.7796 | -34.2254 | -8.3661
4.7857 | -34.2255 | -8.3757
4.7918 | -34.2257 | -8.3848
4.7979 | -34.226 | -8.3941
4.8041 | -34.2261 | -8.4035
4.8102 | -34.2265 | -8.4132
4.8163 | -34.2267 | -8.4228
4.8224 | -34.2268 | -8.4316
4.8285 | -34.2272 | -8.4412
4.8346 | -34.2274 | -8.4511
4.8407 | -34.2276 | -8.4605
4.8468 | -34.2278 | -8.4699
4.8529 | -34.2281 | -8.4794
4.859 | -34.2283 | -8.4886
4.8651 | -34.2286 | -8.4983
4.8712 | -34.2288 | -8.508
4.8773 | -34.2291 | -8.5177
4.8834 | -34.2293 | -8.5275
4.8895 | -34.2296 | -8.5379
4.8956 | -34.2298 | -8.548
4.9017 | -34.2301 | -8.5579
4.9078 | -34.2303 | -8.5674
4.9139 | -34.2305 | -8.5765
4.92 | -34.2307 | -8.5859
4.9261 | -34.231 | -8.5953
4.9322 | -34.2313 | -8.6045
4.9383 | -34.2316 | -8.6137
4.9445 | -34.2319 | -8.6239
4.9506 | -34.2321 | -8.6331
4.9567 | -34.2324 | -8.6426
4.9628 | -34.2327 | -8.6522
4.9689 | -34.2329 | -8.6615
4.975 | -34.2331 | -8.6705
4.9811 | -34.2334 | -8.6808
4.9872 | -34.2335 | -8.6908
4.9933 | -34.2338 | -8.7
4.9994 | -34.2342 | -8.7093
5.0055 | -34.2345 | -8.7189
5.0116 | -34.2348 | -8.7286
5.0177 | -34.2351 | -8.7382
5.0238 | -34.2353 | -8.7483
5.0299 | -34.2357 | -8.7586
5.036 | -34.2359 | -8.7682
5.0421 | -34.2361 | -8.7781
5.0482 | -34.2364 | -8.7883
5.0543 | -34.2367 | -8.7979
5.0604 | -34.237 | -8.8081
5.0665 | -34.2373 | -8.8181
5.0726 | -34.2375 | -8.8279
5.0787 | -34.2378 | -8.8379
5.0848 | -34.2381 | -8.8476
5.091 | -34.2383 | -8.8577
5.0971 | -34.2386 | -8.8681
5.1032 | -34.2389 | -8.8779
5.1093 | -34.2392 | -8.8877
5.1154 | -34.2395 | -8.898
5.1215 | -34.2397 | -8.9083
5.1276 | -34.24 | -8.9182
5.1337 | -34.2403 | -8.928
5.1398 | -34.2405 | -8.9382
5.1459 | -34.2409 | -8.9481
5.152 | -34.2411 | -8.9588
5.1581 | -34.2414 | -8.9692
5.1642 | -34.2416 | -8.979
5.1703 | -34.2419 | -8.9889
5.1764 | -34.2421 | -8.9989
5.1825 | -34.2424 | -9.0095
5.1886 | -34.2426 | -9.0197
5.1947 | -34.2429 | -9.0304
5.2008 | -34.2432 | -9.04
5.2069 | -34.2436 | -9.0508
5.213 | -34.2438 | -9.0606
5.2191 | -34.2442 | -9.0707
5.2252 | -34.2444 | -9.0808
5.2314 | -34.2446 | -9.0911
5.2375 | -34.245 | -9.1013
5.2436 | -34.2453 | -9.1114
5.2497 | -34.2456 | -9.1215
5.2558 | -34.2459 | -9.1319
5.2619 | -34.2461 | -9.1422
5.268 | -34.2464 | -9.1522
5.2741 | -34.2467 | -9.1631
5.2802 | -34.247 | -9.1733
5.2863 | -34.2473 | -9.1834
5.2924 | -34.2475 | -9.1942
5.2985 | -34.2478 | -9.2052
5.3046 | -34.2482 | -9.2162
5.3107 | -34.2486 | -9.2266
5.3168 | -34.249 | -9.2368
5.3229 | -34.2493 | -9.2464
5.329 | -34.2496 | -9.2562
5.3351 | -34.25 | -9.2665
5.3412 | -34.2504 | -9.2766
5.3473 | -34.2507 | -9.2872
5.3534 | -34.251 | -9.2977
5.3595 | -34.2513 | -9.3082
5.3656 | -34.2516 | -9.319
5.3717 | -34.2519 | -9.3289
5.3779 | -34.2522 | -9.3391
5.384 | -34.2525 | -9.3496
5.3901 | -34.2528 | -9.3597
5.3962 | -34.253 | -9.3705
5.4023 | -34.2534 | -9.3807
5.4084 | -34.2536 | -9.3909
5.4145 | -34.2539 | -9.402
5.4206 | -34.2541 | -9.413
5.4267 | -34.2545 | -9.4236
5.4328 | -34.2548 | -9.4335
5.4389 | -34.2551 | -9.4439
5.445 | -34.2554 | -9.4545
5.4511 | -34.2558 | -9.4648
5.4572 | -34.2561 | -9.4756
5.4633 | -34.2564 | -9.4857
5.4694 | -34.2567 | -9.4959
5.4755 | -34.257 | -9.5069
5.4816 | -34.2573 | -9.5171
5.4877 | -34.2576 | -9.5277
5.4938 | -34.2579 | -9.5382
5.4999 | -34.2581 | -9.5488
5.506 | -34.2584 | -9.5593
5.5121 | -34.2587 | -9.5703
5.5183 | -34.259 | -9.5805
5.5244 | -34.2593 | -9.5908
5.5305 | -34.2596 | -9.6015
5.5366 | -34.2599 | -9.6121
5.5427 | -34.2602 | -9.6229
5.5488 | -34.2605 | -9.6337
5.5549 | -34.2608 | -9.6441
5.561 | -34.2611 | -9.6548
5.5671 | -34.2614 | -9.6652
5.5732 | -34.2617 | -9.6756
5.5793 | -34.262 | -9.6864
5.5854 | -34.2623 | -9.6971
5.5915 | -34.2625 | -9.7074
5.5976 | -34.2629 | -9.7181
5.6037 | -34.2632 | -9.7285
5.6098 | -34.2635 | -9.7388
5.6159 | -34.2638 | -9.7497
5.622 | -34.2641 | -9.7598
5.6281 | -34.2644 | -9.7702
5.6342 | -34.2647 | -9.7804
5.6403 | -34.265 | -9.7903
5.6464 | -34.2654 | -9.8007
5.6525 | -34.2656 | -9.8107
5.6586 | -34.266 | -9.8208
5.6648 | -34.2662 | -9.8317
5.6709 | -34.2665 | -9.8422
5.677 | -34.2668 | -9.8533
5.6831 | -34.2671 | -9.8633
5.6892 | -34.2674 | -9.8736
5.6953 | -34.2677 | -9.8839
5.7014 | -34.268 | -9.8948
5.7075 | -34.2684 | -9.9056
5.7136 | -34.2687 | -9.9159
5.7197 | -34.269 | -9.9263
5.7258 | -34.2693 | -9.9369
5.7319 | -34.2696 | -9.9472
5.738 | -34.2699 | -9.9576
5.7441 | -34.2702 | -9.9681
5.7502 | -34.2705 | -9.9784
5.7563 | -34.2708 | -9.9885
5.7624 | -34.271 | -9.9994
5.7685 | -34.2713 | -10.0096
5.7746 | -34.2715 | -10.0195
5.7807 | -34.2719 | -10.0295
5.7868 | -34.2722 | -10.0396
5.7929 | -34.2725 | -10.0501
5.799 | -34.2728 | -10.0606
5.8052 | -34.2731 | -10.0713
5.8113 | -34.2733 | -10.0819
5.8174 | -34.2736 | -10.0926
5.8235 | -34.2739 | -10.1029
5.8296 | -34.2742 | -10.1137
5.8357 | -34.2745 | -10.1243
5.8418 | -34.2748 | -10.1347
5.8479 | -34.275 | -10.1453
5.854 | -34.2753 | -10.1557
5.8601 | -34.2756 | -10.1661
5.8662 | -34.2759 | -10.1766
5.8723 | -34.2762 | -10.1866
5.8784 | -34.2765 | -10.1971
5.8845 | -34.2768 | -10.2078
5.8906 | -34.2771 | -10.218
5.8967 | -34.2774 | -10.2283
5.9028 | -34.2777 | -10.2389
5.9089 | -34.2778 | -10.2487
5.915 | -34.2781 | -10.2594
5.9211 | -34.2785 | -10.2699
5.9272 | -34.2787 | -10.2806
5.9333 | -34.279 | -10.2905
5.9394 | -34.2793 | -10.3006
5.9455 | -34.2796 | -10.3108
5.9517 | -34.2799 | -10.3213
5.9578 | -34.2802 | -10.3317
5.9639 | -34.2805 | -10.3422
5.97 | -34.2807 | -10.3524
5.9761 | -34.281 | -10.3626
5.9822 | -34.2814 | -10.3729
5.9883 | -34.2817 | -10.383
5.9944 | -34.282 | -10.3932
6.0005 | -34.2823 | -10.4035
6.0066 | -34.2825 | -10.4138
6.0127 | -34.2828 | -10.4237
6.0188 | -34.2831 | -10.434
6.0249 | -34.2834 | -10.4442
6.031 | -34.2837 | -10.4546
6.0371 | -34.2841 | -10.4651
6.0432 | -34.2844 | -10.4756
6.0493 | -34.2848 | -10.4858
6.0554 | -34.2851 | -10.4962
6.0615 | -34.2853 | -10.507
6.0676 | -34.2856 | -10.5176
6.0737 | -34.2859 | -10.5279
6.0798 | -34.2862 | -10.5381
6.0859 | -34.2865 | -10.5483
6.0921 | -34.2868 | -10.5587
6.0982 | -34.2872 | -10.5693
6.1043 | -34.2875 | -10.5795
6.1104 | -34.2878 | -10.59
6.1165 | -34.2881 | -10.6
6.1226 | -34.2885 | -10.6098
6.1287 | -34.2888 | -10.6203
6.1348 | -34.2892 | -10.6306
6.1409 | -34.2894 | -10.6409
6.147 | -34.2898 | -10.6511
6.1531 | -34.2901 | -10.6614
6.1592 | -34.2904 | -10.6718
6.1653 | -34.2907 | -10.6821
6.1714 | -34.2911 | -10.6923
6.1775 | -34.2914 | -10.7027
6.1836 | -34.2917 | -10.7128
6.1897 | -34.292 | -10.723
6.1958 | -34.2923 | -10.7331
6.2019 | -34.2926 | -10.7435
6.208 | -34.2929 | -10.7542
6.2141 | -34.2932 | -10.7647
6.2202 | -34.2935 | -10.7753
6.2263 | -34.2939 | -10.7854
6.2325 | -34.2942 | -10.7956
6.2386 | -34.2946 | -10.8059
6.2447 | -34.2948 | -10.816
6.2508 | -34.2951 | -10.8262
6.2569 | -34.2954 | -10.8368
6.263 | -34.2958 | -10.8468
6.2691 | -34.2962 | -10.8571
6.2752 | -34.2966 | -10.8672
6.2813 | -34.2969 | -10.8775
6.2874 | -34.2972 | -10.8879
6.2935 | -34.2976 | -10.8979
6.2996 | -34.2979 | -10.9084
6.3057 | -34.2983 | -10.9184
6.3118 | -34.2987 | -10.9286
6.3179 | -34.299 | -10.9388
6.324 | -34.2993 | -10.9492
6.3301 | -34.2997 | -10.9592
6.3362 | -34.3 | -10.9696
6.3423 | -34.3004 | -10.9796
6.3484 | -34.3008 | -10.9899
6.3545 | -34.3012 | -11
6.3606 | -34.3015 | -11.0104
6.3667 | -34.302 | -11.0205
6.3728 | -34.3023 | -11.0307
6.379 | -34.3026 | -11.0409
6.3851 | -34.303 | -11.0515
6.3912 | -34.3034 | -11.0616
6.3973 | -34.3038 | -11.0714
6.4034 | -34.3042 | -11.0818
6.4095 | -34.3045 | -11.0923
6.4156 | -34.3048 | -11.1022
6.4217 | -34.3052 | -11.1125
6.4278 | -34.3055 | -11.1225
6.4339 | -34.3059 | -11.1327
6.44 | -34.3062 | -11.1426
6.4461 | -34.3066 | -11.1531
6.4522 | -34.307 | -11.163
6.4583 | -34.3073 | -11.1733
6.4644 | -34.3076 | -11.1837
6.4705 | -34.3081 | -11.1944
6.4766 | -34.3084 | -11.2049
6.4827 | -34.3088 | -11.2156
6.4888 | -34.3091 | -11.2259
6.4949 | -34.3095 | -11.2361
6.501 | -34.3098 | -11.2464
6.5071 | -34.3102 | -11.2569
6.5132 | -34.3106 | -11.267
6.5194 | -34.311 | -11.277
6.5255 | -34.3113 | -11.2871
6.5316 | -34.3117 | -11.2974
6.5377 | -34.312 | -11.3078
6.5438 | -34.3124 | -11.3181
6.5499 | -34.3128 | -11.3282
6.556 | -34.3132 | -11.3384
6.5621 | -34.3135 | -11.3485
6.5682 | -34.3139 | -11.3584
6.5743 | -34.3142 | -11.3693
6.5804 | -34.3145 | -11.3794
6.5865 | -34.3148 | -11.3898
6.5926 | -34.3152 | -11.4001
6.5987 | -34.3156 | -11.4103
6.6048 | -34.3159 | -11.4211
6.6109 | -34.3163 | -11.431
6.617 | -34.3167 | -11.442
6.6231 | -34.3169 | -11.4524
6.6292 | -34.3174 | -11.4625
6.6353 | -34.3178 | -11.4733
6.6414 | -34.3182 | -11.484
6.6475 | -34.3185 | -11.4942
6.6536 | -34.3189 | -11.5041
6.6597 | -34.3193 | -11.5145
6.6659 | -34.3197 | -11.5245
6.672 | -34.3201 | -11.5349
6.6781 | -34.3204 | -11.5449
6.6842 | -34.3208 | -11.5554
6.6903 | -34.3211 | -11.5656
6.6964 | -34.3215 | -11.5759
6.7025 | -34.3218 | -11.5864
6.7086 | -34.3222 | -11.5964
6.7147 | -34.3226 | -11.6066
6.7208 | -34.323 | -11.6173
6.7269 | -34.3233 | -11.6272
6.733 | -34.3236 | -11.637
6.7391 | -34.324 | -11.6474
6.7452 | -34.3245 | -11.6573
6.7513 | -34.3248 | -11.6674
6.7574 | -34.3252 | -11.6778
6.7635 | -34.3256 | -11.6878
6.7696 | -34.326 | -11.6982
6.7757 | -34.3264 | -11.7086
6.7818 | -34.3268 | -11.7189
6.7879 | -34.3271 | -11.7292
6.794 | -34.3275 | -11.7395
6.8001 | -34.3279 | -11.7496
6.8063 | -34.3283 | -11.76
6.8124 | -34.3286 | -11.7701
6.8185 | -34.329 | -11.78
6.8246 | -34.3294 | -11.7903
6.8307 | -34.3298 | -11.8006
6.8368 | -34.3302 | -11.8107
6.8429 | -34.3306 | -11.8207
6.849 | -34.3309 | -11.8309
6.8551 | -34.3313 | -11.8411
6.8612 | -34.3317 | -11.8511
6.8673 | -34.332 | -11.8612
6.8734 | -34.3324 | -11.8718
6.8795 | -34.3328 | -11.8823
6.8856 | -34.3332 | -11.892
6.8917 | -34.3336 | -11.9016
6.8978 | -34.3339 | -11.912
6.9039 | -34.3343 | -11.9224
6.91 | -34.3346 | -11.9325
6.9161 | -34.335 | -11.9427
6.9222 | -34.3353 | -11.953
6.9283 | -34.3356 | -11.963
6.9344 | -34.336 | -11.9729
6.9405 | -34.3363 | -11.983
6.9466 | -34.3366 | -11.993
6.9528 | -34.337 | -12.0029
6.9589 | -34.3373 | -12.0127
6.965 | -34.3377 | -12.0226
6.9711 | -34.3381 | -12.0325
6.9772 | -34.3384 | -12.0424
6.9833 | -34.3387 | -12.0528
6.9894 | -34.3391 | -12.063
6.9955 | -34.3394 | -12.0733
7.0016 | -34.3398 | -12.0838
7.0077 | -34.3402 | -12.0942
7.0138 | -34.3405 | -12.1047
7.0199 | -34.3409 | -12.1148
7.026 | -34.3411 | -12.1248
7.0321 | -34.3415 | -12.1351
7.0382 | -34.3418 | -12.1454
7.0443 | -34.3422 | -12.1556
7.0504 | -34.3425 | -12.166
7.0565 | -34.3428 | -12.1763
7.0626 | -34.3431 | -12.1866
7.0687 | -34.3435 | -12.1967
7.0748 | -34.3438 | -12.207
7.0809 | -34.3441 | -12.2173
7.087 | -34.3445 | -12.2273
7.0932 | -34.3448 | -12.2374
7.0993 | -34.3452 | -12.2476
7.1054 | -34.3455 | -12.2575
7.1115 | -34.3457 | -12.2678
7.1176 | -34.346 | -12.2779
7.1237 | -34.3463 | -12.2881
7.1298 | -34.3466 | -12.2979
7.1359 | -34.3469 | -12.3077
7.142 | -34.3473 | -12.3182
7.1481 | -34.3476 | -12.3282
7.1542 | -34.3479 | -12.3381
7.1603 | -34.3482 | -12.3481
7.1664 | -34.3486 | -12.3582
7.1725 | -34.3489 | -12.368
7.1786 | -34.3492 | -12.3781
7.1847 | -34.3496 | -12.388
7.1908 | -34.3499 | -12.398
7.1969 | -34.3503 | -12.4079
7.203 | -34.3506 | -12.4183
7.2091 | -34.3509 | -12.4282
7.2152 | -34.3512 | -12.4382
7.2213 | -34.3515 | -12.4483
7.2274 | -34.3518 | -12.4584
7.2335 | -34.3521 | -12.4686
7.2397 | -34.3524 | -12.4787
7.2458 | -34.3527 | -12.4885
7.2519 | -34.353 | -12.4984
7.258 | -34.3533 | -12.5084
7.2641 | -34.3537 | -12.5184
7.2702 | -34.354 | -12.5285
7.2763 | -34.3543 | -12.5382
7.2824 | -34.3547 | -12.5479
7.2885 | -34.3551 | -12.5578
7.2946 | -34.3554 | -12.568
7.3007 | -34.3557 | -12.578
7.3068 | -34.356 | -12.588
7.3129 | -34.3563 | -12.5976
7.319 | -34.3567 | -12.6077
7.3251 | -34.357 | -12.6176
7.3312 | -34.3573 | -12.6272
7.3373 | -34.3577 | -12.6371
7.3434 | -34.358 | -12.647
7.3495 | -34.3583 | -12.6566
7.3556 | -34.3586 | -12.6668
7.3617 | -34.3589 | -12.6769
7.3678 | -34.3592 | -12.6871
7.3739 | -34.3596 | -12.6965
7.3801 | -34.3599 | -12.706
7.3862 | -34.3602 | -12.7158
7.3923 | -34.3605 | -12.7255
7.3984 | -34.3608 | -12.7356
7.4045 | -34.3611 | -12.7454
7.4106 | -34.3614 | -12.7552
7.4167 | -34.3617 | -12.7653
7.4228 | -34.3621 | -12.7748
7.4289 | -34.3624 | -12.7845
7.435 | -34.3628 | -12.7944
7.4411 | -34.3631 | -12.8039
7.4472 | -34.3634 | -12.8139
7.4533 | -34.3638 | -12.824
7.4594 | -34.364 | -12.8339
7.4655 | -34.3644 | -12.8439
7.4716 | -34.3646 | -12.8536
7.4777 | -34.365 | -12.8631
7.4838 | -34.3653 | -12.8734
7.4899 | -34.3657 | -12.8834
7.496 | -34.3661 | -12.8936
7.5021 | -34.3664 | -12.9034
7.5082 | -34.3667 | -12.9133
7.5143 | -34.367 | -12.923
7.5204 | -34.3673 | -12.9324
7.5266 | -34.3676 | -12.9425
7.5327 | -34.368 | -12.9523
7.5388 | -34.3683 | -12.9618
7.5449 | -34.3685 | -12.9715
7.551 | -34.3688 | -12.9817
7.5571 | -34.3691 | -12.9913
7.5632 | -34.3695 | -13.0009
7.5693 | -34.3698 | -13.011
7.5754 | -34.3701 | -13.0211
7.5815 | -34.3705 | -13.0305
7.5876 | -34.3707 | -13.0399
7.5937 | -34.371 | -13.0494
7.5998 | -34.3713 | -13.0593
7.6059 | -34.3715 | -13.0687
7.612 | -34.3719 | -13.0785
7.6181 | -34.3722 | -13.0883
7.6242 | -34.3726 | -13.0978
7.6303 | -34.3729 | -13.1071
7.6364 | -34.3733 | -13.117
7.6425 | -34.3736 | -13.127
7.6486 | -34.3739 | -13.1368
7.6547 | -34.3742 | -13.1468
7.6608 | -34.3746 | -13.1566
7.667 | -34.3749 | -13.1657
7.6731 | -34.3752 | -13.1755
7.6792 | -34.3755 | -13.1851
7.6853 | -34.3758 | -13.1945
7.6914 | -34.376 | -13.2043
7.6975 | -34.3763 | -13.2141
7.7036 | -34.3766 | -13.2238
7.7097 | -34.377 | -13.2332
7.7158 | -34.3773 | -13.2427
7.7219 | -34.3775 | -13.2526
7.728 | -34.3778 | -13.2623
7.7341 | -34.3781 | -13.272
7.7402 | -34.3785 | -13.2814
7.7463 | -34.3787 | -13.2911
7.7524 | -34.379 | -13.3008
7.7585 | -34.3793 | -13.3106
7.7646 | -34.3796 | -13.3201
7.7707 | -34.3799 | -13.3294
7.7768 | -34.3802 | -13.3388
7.7829 | -34.3805 | -13.3482
7.789 | -34.3808 | -13.3573
7.7951 | -34.3811 | -13.3673
7.8012 | -34.3814 | -13.3773
7.8073 | -34.3817 | -13.3864
7.8135 | -34.3819 | -13.3962
7.8196 | -34.3822 | -13.4055
7.8257 | -34.3826 | -13.4152
7.8318 | -34.383 | -13.4253
7.8379 | -34.3832 | -13.4345
7.844 | -34.3835 | -13.4438
7.8501 | -34.3837 | -13.4535
7.8562 | -34.384 | -13.4634
7.8623 | -34.3843 | -13.4731
7.8684 | -34.3846 | -13.4824
7.8745 | -34.3849 | -13.4918
7.8806 | -34.3852 | -13.5007
7.8867 | -34.3855 | -13.5104
7.8928 | -34.3858 | -13.5201
7.8989 | -34.3862 | -13.5292
7.905 | -34.3865 | -13.5391
7.9111 | -34.3869 | -13.5486
7.9172 | -34.3872 | -13.5579
7.9233 | -34.3875 | -13.5675
7.9294 | -34.3878 | -13.5771
7.9355 | -34.3881 | -13.5865
7.9416 | -34.3884 | -13.5959
7.9477 | -34.3887 | -13.6052
7.9539 | -34.389 | -13.6145
7.96 | -34.3893 | -13.6242
7.9661 | -34.3896 | -13.6338
7.9722 | -34.3899 | -13.6436
7.9783 | -34.3902 | -13.6529
7.9844 | -34.3906 | -13.6628
7.9905 | -34.3909 | -13.672
7.9966 | -34.3911 | -13.6814
8.0027 | -34.3915 | -13.6913
8.0088 | -34.3918 | -13.7004
8.0149 | -34.3921 | -13.7096
8.021 | -34.3924 | -13.7191
8.0271 | -34.3927 | -13.7287
8.0332 | -34.393 | -13.7383
8.0393 | -34.3933 | -13.7478
8.0454 | -34.3936 | -13.7572
8.0515 | -34.3939 | -13.7664
8.0576 | -34.3942 | -13.7759
8.0637 | -34.3945 | -13.7852
8.0698 | -34.3948 | -13.7944
8.0759 | -34.3952 | -13.8038
8.082 | -34.3955 | -13.8135
8.0881 | -34.3958 | -13.8228
8.0942 | -34.396 | -13.8325
8.1004 | -34.3963 | -13.8422
8.1065 | -34.3967 | -13.8518
8.1126 | -34.397 | -13.8612
8.1187 | -34.3974 | -13.8707
8.1248 | -34.3977 | -13.8806
8.1309 | -34.398 | -13.8902
8.137 | -34.3982 | -13.8999
8.1431 | -34.3986 | -13.9092
8.1492 | -34.3989 | -13.9189
8.1553 | -34.3992 | -13.9283
8.1614 | -34.3996 | -13.9377
8.1675 | -34.3999 | -13.9471
8.1736 | -34.4002 | -13.9567
8.1797 | -34.4005 | -13.9662
8.1858 | -34.4008 | -13.9756
8.1919 | -34.4011 | -13.9849
8.198 | -34.4014 | -13.9943
8.2041 | -34.4017 | -14.004
8.2102 | -34.4021 | -14.0134
8.2163 | -34.4024 | -14.0231
8.2224 | -34.4027 | -14.0328
8.2285 | -34.4031 | -14.0424
8.2346 | -34.4034 | -14.0519
8.2408 | -34.4037 | -14.0616
8.2469 | -34.404 | -14.0708
8.253 | -34.4042 | -14.0804
8.2591 | -34.4046 | -14.0894
8.2652 | -34.405 | -14.0985
8.2713 | -34.4053 | -14.1079
8.2774 | -34.4056 | -14.1173
8.2835 | -34.406 | -14.1268
8.2896 | -34.4063 | -14.136
8.2957 | -34.4066 | -14.1454
8.3018 | -34.4069 | -14.1547
8.3079 | -34.4073 | -14.1642
8.314 | -34.4076 | -14.1735
8.3201 | -34.4079 | -14.1829
8.3262 | -34.4082 | -14.1922
8.3323 | -34.4086 | -14.2016
8.3384 | -34.4089 | -14.2113
8.3445 | -34.4092 | -14.2205
8.3506 | -34.4095 | -14.23
8.3567 | -34.4098 | -14.2396
8.3628 | -34.4101 | -14.249
8.3689 | -34.4105 | -14.2586
8.375 | -34.4108 | -14.268
8.3812 | -34.4112 | -14.2777
8.3873 | -34.4114 | -14.2871
8.3934 | -34.4118 | -14.2964
8.3995 | -34.4121 | -14.3062
8.4056 | -34.4125 | -14.3157
8.4117 | -34.4128 | -14.3253
8.4178 | -34.4132 | -14.3343
8.4239 | -34.4135 | -14.3437
8.43 | -34.4139 | -14.3532
8.4361 | -34.4143 | -14.3629
8.4422 | -34.4145 | -14.3723
8.4483 | -34.4148 | -14.3818
8.4544 | -34.4151 | -14.3915
8.4605 | -34.4154 | -14.4011
8.4666 | -34.4158 | -14.4104
8.4727 | -34.4161 | -14.4198
8.4788 | -34.4165 | -14.429
8.4849 | -34.4168 | -14.4388
8.491 | -34.4171 | -14.4481
8.4971 | -34.4174 | -14.4578
8.5032 | -34.4177 | -14.4675
8.5093 | -34.4181 | -14.477
8.5154 | -34.4184 | -14.4861
8.5215 | -34.4186 | -14.4958
8.5277 | -34.419 | -14.5053
8.5338 | -34.4193 | -14.5146
8.5399 | -34.4196 | -14.5243
8.546 | -34.42 | -14.5335
8.5521 | -34.4203 | -14.5431
8.5582 | -34.4207 | -14.5527
8.5643 | -34.421 | -14.5623
8.5704 | -34.4213 | -14.5718
8.5765 | -34.4216 | -14.581
8.5826 | -34.422 | -14.5908
8.5887 | -34.4223 | -14.6001
8.5948 | -34.4225 | -14.6091
8.6009 | -34.4229 | -14.6188
8.607 | -34.4232 | -14.6283
8.6131 | -34.4235 | -14.638
8.6192 | -34.4238 | -14.6472
8.6253 | -34.4241 | -14.6569
8.6314 | -34.4245 | -14.6664
8.6375 | -34.4248 | -14.6761
8.6436 | -34.4252 | -14.6855
8.6497 | -34.4256 | -14.6951
8.6558 | -34.4259 | -14.7044
8.6619 | -34.4263 | -14.714
8.6681 | -34.4265 | -14.7232
8.6742 | -34.4268 | -14.7328
8.6803 | -34.4272 | -14.7427
8.6864 | -34.4276 | -14.7527
8.6925 | -34.4279 | -14.7625
8.6986 | -34.4282 | -14.7722
8.7047 | -34.4286 | -14.7816
8.7108 | -34.4289 | -14.7911
8.7169 | -34.4293 | -14.8003
8.723 | -34.4297 | -14.81
8.7291 | -34.43 | -14.8196
8.7352 | -34.4304 | -14.8291
8.7413 | -34.4307 | -14.8384
8.7474 | -34.4311 | -14.8479
8.7535 | -34.4314 | -14.8574
8.7596 | -34.4318 | -14.8673
8.7657 | -34.4321 | -14.8767
8.7718 | -34.4325 | -14.8858
8.7779 | -34.4328 | -14.8958
8.784 | -34.4331 | -14.905
8.7901 | -34.4333 | -14.9148
8.7962 | -34.4337 | -14.9247
8.8023 | -34.4342 | -14.934
8.8084 | -34.4345 | -14.9437
8.8146 | -34.4348 | -14.9531
8.8207 | -34.4352 | -14.9628
8.8268 | -34.4355 | -14.9724
8.8329 | -34.4359 | -14.982
8.839 | -34.4362 | -14.9915
8.8451 | -34.4366 | -15.0011
8.8512 | -34.437 | -15.0108
8.8573 | -34.4373 | -15.0201
8.8634 | -34.4377 | -15.0297
8.8695 | -34.438 | -15.0391
8.8756 | -34.4384 | -15.0484
8.8817 | -34.4387 | -15.0583
8.8878 | -34.439 | -15.0678
8.8939 | -34.4393 | -15.077
8.9 | -34.4397 | -15.0861
8.9061 | -34.4401 | -15.0958
8.9122 | -34.4404 | -15.1051
8.9183 | -34.4407 | -15.1144
8.9244 | -34.4411 | -15.1234
8.9305 | -34.4414 | -15.1331
8.9366 | -34.4417 | -15.1424
8.9427 | -34.4421 | -15.1519
8.9488 | -34.4425 | -15.1614
8.955 | -34.4428 | -15.171
8.9611 | -34.4432 | -15.1803
8.9672 | -34.4435 | -15.19
8.9733 | -34.4438 | -15.1997
8.9794 | -34.4442 | -15.2094
8.9855 | -34.4445 | -15.2187
8.9916 | -34.4449 | -15.2279
8.9977 | -34.4452 | -15.2373
9.0038 | -34.4456 | -15.2469
9.0099 | -34.4459 | -15.2565
9.016 | -34.4464 | -15.266
9.0221 | -34.4467 | -15.276
9.0282 | -34.447 | -15.2859
9.0343 | -34.4474 | -15.2957
9.0404 | -34.4478 | -15.3049
9.0465 | -34.4482 | -15.3146
9.0526 | -34.4485 | -15.324
9.0587 | -34.4488 | -15.3333
9.0648 | -34.4492 | -15.3429
9.0709 | -34.4495 | -15.3522
9.077 | -34.4499 | -15.362
9.0831 | -34.4502 | -15.3714
9.0892 | -34.4506 | -15.3809
9.0953 | -34.4508 | -15.3902
9.1015 | -34.4512 | -15.3998
9.1076 | -34.4516 | -15.4094
9.1137 | -34.452 | -15.4188
9.1198 | -34.4523 | -15.4279
9.1259 | -34.4527 | -15.4375
9.132 | -34.4531 | -15.4471
9.1381 | -34.4534 | -15.4565
9.1442 | -34.4538 | -15.4657
9.1503 | -34.4541 | -15.4753
9.1564 | -34.4545 | -15.4849
9.1625 | -34.4549 | -15.4944
9.1686 | -34.4552 | -15.5039
9.1747 | -34.4556 | -15.5135
9.1808 | -34.4559 | -15.523
9.1869 | -34.4562 | -15.5327
9.193 | -34.4566 | -15.5423
9.1991 | -34.4569 | -15.5515
9.2052 | -34.4573 | -15.5609
9.2113 | -34.4577 | -15.5705
9.2174 | -34.458 | -15.5805
9.2235 | -34.4584 | -15.5898
9.2296 | -34.4588 | -15.5993
9.2357 | -34.4591 | -15.6086
9.2419 | -34.4595 | -15.618
9.248 | -34.4598 | -15.6275
9.2541 | -34.4602 | -15.6371
9.2602 | -34.4605 | -15.6465
9.2663 | -34.4609 | -15.6561
9.2724 | -34.4612 | -15.6657
9.2785 | -34.4616 | -15.6752
9.2846 | -34.4619 | -15.6848
9.2907 | -34.4623 | -15.6942
9.2968 | -34.4626 | -15.7036
9.3029 | -34.4629 | -15.713
9.309 | -34.4633 | -15.7225
9.3151 | -34.4637 | -15.7321
9.3212 | -34.464 | -15.7418
9.3273 | -34.4644 | -15.7512
9.3334 | -34.4648 | -15.7606
9.3395 | -34.4653 | -15.7702
9.3456 | -34.4657 | -15.7796
9.3517 | -34.466 | -15.7886
9.3578 | -34.4664 | -15.7977
9.3639 | -34.4668 | -15.8076
9.37 | -34.4672 | -15.817
9.3761 | -34.4675 | -15.8261
9.3822 | -34.4678 | -15.8354
9.3884 | -34.4682 | -15.845
9.3945 | -34.4686 | -15.8544
9.4006 | -34.4689 | -15.8639
9.4067 | -34.4693 | -15.8735
9.4128 | -34.4696 | -15.8828
9.4189 | -34.47 | -15.8921
9.425 | -34.4704 | -15.9014
9.4311 | -34.4707 | -15.9108
9.4372 | -34.4711 | -15.9202
9.4433 | -34.4714 | -15.9298
9.4494 | -34.4718 | -15.9394
9.4555 | -34.4721 | -15.9486
9.4616 | -34.4725 | -15.9579
9.4677 | -34.4729 | -15.9672
9.4738 | -34.4733 | -15.977
9.4799 | -34.4736 | -15.9863
9.486 | -34.4739 | -15.996
9.4921 | -34.4743 | -16.006
9.4982 | -34.4747 | -16.015
9.5043 | -34.475 | -16.0242
9.5104 | -34.4754 | -16.0337
9.5165 | -34.4758 | -16.043
9.5226 | -34.4762 | -16.0525
9.5288 | -34.4765 | -16.0619
9.5349 | -34.4769 | -16.0713
9.541 | -34.4773 | -16.081
9.5471 | -34.4776 | -16.0906
9.5532 | -34.478 | -16.0999
9.5593 | -34.4784 | -16.109
9.5654 | -34.4788 | -16.1185
9.5715 | -34.4792 | -16.1283
9.5776 | -34.4796 | -16.1377
9.5837 | -34.4799 | -16.147
9.5898 | -34.4803 | -16.1566
9.5959 | -34.4807 | -16.1657
9.602 | -34.481 | -16.1754
9.6081 | -34.4814 | -16.1849
9.6142 | -34.4818 | -16.1943
9.6203 | -34.4821 | -16.2038
9.6264 | -34.4825 | -16.213
9.6325 | -34.4828 | -16.2225
9.6386 | -34.4832 | -16.2319
9.6447 | -34.4836 | -16.2412
9.6508 | -34.4839 | -16.2508
9.6569 | -34.4843 | -16.2602
9.663 | -34.4847 | -16.2695
9.6691 | -34.4851 | -16.2788
9.6753 | -34.4855 | -16.2883
9.6814 | -34.4859 | -16.2977
9.6875 | -34.4863 | -16.3072
9.6936 | -34.4867 | -16.3167
9.6997 | -34.487 | -16.3261
9.7058 | -34.4874 | -16.3353
9.7119 | -34.4877 | -16.3445
9.718 | -34.4881 | -16.354
9.7241 | -34.4884 | -16.3636
9.7302 | -34.4888 | -16.3731
9.7363 | -34.4892 | -16.3823
9.7424 | -34.4896 | -16.3916
9.7485 | -34.49 | -16.4009
9.7546 | -34.4904 | -16.4103
9.7607 | -34.4907 | -16.4197
9.7668 | -34.4911 | -16.4294
9.7729 | -34.4915 | -16.4386
9.779 | -34.4919 | -16.4481
9.7851 | -34.4922 | -16.4574
9.7912 | -34.4926 | -16.4674
9.7973 | -34.4929 | -16.4768
9.8034 | -34.4933 | -16.4866
9.8095 | -34.4936 | -16.4956
9.8157 | -34.494 | -16.5049
9.8218 | -34.4944 | -16.5137
9.8279 | -34.4948 | -16.5234
9.834 | -34.4952 | -16.5326
9.8401 | -34.4956 | -16.5418
9.8462 | -34.496 | -16.551
9.8523 | -34.4964 | -16.5604
9.8584 | -34.4967 | -16.57
9.8645 | -34.4971 | -16.5792
9.8706 | -34.4975 | -16.5884
9.8767 | -34.4979 | -16.5975
9.8828 | -34.4984 | -16.6066
9.8889 | -34.4987 | -16.6164
9.895 | -34.4991 | -16.6254
9.9011 | -34.4994 | -16.6343
9.9072 | -34.4998 | -16.6434
9.9133 | -34.5003 | -16.6527
9.9194 | -34.5007 | -16.6618
9.9255 | -34.501 | -16.6714
9.9316 | -34.5015 | -16.6805
9.9377 | -34.5019 | -16.6897
9.9438 | -34.5023 | -16.6989
9.9499 | -34.5027 | -16.7081
9.956 | -34.5031 | -16.7171
9.9622 | -34.5035 | -16.7262
9.9683 | -34.5039 | -16.7355
9.9744 | -34.5043 | -16.7446
9.9805 | -34.5046 | -16.7539
9.9866 | -34.5051 | -16.763
9.9927 | -34.5055 | -16.7726
9.9988 | -34.5059 | -16.7815
10.0049 | -34.5062 | -16.7907
10.011 | -34.5067 | -16.7999
10.0171 | -34.507 | -16.8093
10.0232 | -34.5075 | -16.8184
10.0293 | -34.5079 | -16.8273
10.0354 | -34.5083 | -16.8362
10.0415 | -34.5087 | -16.8453
10.0476 | -34.5091 | -16.8548
10.0537 | -34.5094 | -16.8642
10.0598 | -34.5099 | -16.8736
10.0659 | -34.5102 | -16.8828
10.072 | -34.5107 | -16.8922
10.0781 | -34.5111 | -16.9015
10.0842 | -34.5114 | -16.9109
10.0903 | -34.5118 | -16.9204
10.0964 | -34.5123 | -16.9296
10.1026 | -34.5128 | -16.9388
10.1087 | -34.5132 | -16.9482
10.1148 | -34.5136 | -16.9575
10.1209 | -34.514 | -16.9667
10.127 | -34.5145 | -16.9758
10.1331 | -34.5148 | -16.9851
10.1392 | -34.5153 | -16.9944
10.1453 | -34.5156 | -17.0038
10.1514 | -34.516 | -17.0127
10.1575 | -34.5164 | -17.0218
10.1636 | -34.5169 | -17.0305
10.1697 | -34.5173 | -17.0396
10.1758 | -34.5177 | -17.0489
10.1819 | -34.5181 | -17.0583
10.188 | -34.5185 | -17.0673
10.1941 | -34.5189 | -17.0766
10.2002 | -34.5193 | -17.0852
10.2063 | -34.5197 | -17.0947
10.2124 | -34.5201 | -17.1038
10.2185 | -34.5205 | -17.1128
10.2246 | -34.5209 | -17.1221
10.2307 | -34.5212 | -17.1314
10.2368 | -34.5217 | -17.14
10.2429 | -34.5221 | -17.1489
10.2491 | -34.5225 | -17.1578
10.2552 | -34.523 | -17.1673
10.2613 | -34.5234 | -17.1766
10.2674 | -34.5238 | -17.1858
10.2735 | -34.5241 | -17.1949
10.2796 | -34.5245 | -17.2042
10.2857 | -34.5249 | -17.2132
10.2918 | -34.5253 | -17.2223
10.2979 | -34.5258 | -17.2317
10.304 | -34.5262 | -17.2411
10.3101 | -34.5267 | -17.2503
10.3162 | -34.5271 | -17.2593
10.3223 | -34.5274 | -17.2687
10.3284 | -34.5279 | -17.2777
10.3345 | -34.5283 | -17.2868
10.3406 | -34.5287 | -17.2959
10.3467 | -34.5291 | -17.3052
10.3528 | -34.5295 | -17.3142
10.3589 | -34.5299 | -17.3236
10.365 | -34.5303 | -17.3328
10.3711 | -34.5307 | -17.342
10.3772 | -34.531 | -17.3511
10.3833 | -34.5314 | -17.3601
10.3895 | -34.5318 | -17.3691
10.3956 | -34.5323 | -17.378
10.4017 | -34.5326 | -17.3872
10.4078 | -34.533 | -17.3965
10.4139 | -34.5335 | -17.4054
10.42 | -34.5339 | -17.4144
10.4261 | -34.5344 | -17.4235
10.4322 | -34.5348 | -17.4325
10.4383 | -34.5352 | -17.4418
10.4444 | -34.5356 | -17.4507
10.4505 | -34.5361 | -17.4597
10.4566 | -34.5365 | -17.4689
10.4627 | -34.5369 | -17.4782
10.4688 | -34.5373 | -17.487
10.4749 | -34.5377 | -17.4959
10.481 | -34.5382 | -17.5049
10.4871 | -34.5385 | -17.5142
10.4932 | -34.539 | -17.5234
10.4993 | -34.5395 | -17.5323
10.5054 | -34.5399 | -17.5415
10.5115 | -34.5403 | -17.5505
10.5176 | -34.5407 | -17.5594
10.5237 | -34.5411 | -17.5681
10.5298 | -34.5415 | -17.5771
10.536 | -34.542 | -17.5864
10.5421 | -34.5423 | -17.5951
10.5482 | -34.5428 | -17.6045
10.5543 | -34.5432 | -17.6136
10.5604 | -34.5436 | -17.6226
10.5665 | -34.544 | -17.6316
10.5726 | -34.5444 | -17.6405
10.5787 | -34.5449 | -17.6498
10.5848 | -34.5453 | -17.6588
10.5909 | -34.5457 | -17.6678
10.597 | -34.5461 | -17.6768
10.6031 | -34.5466 | -17.6858
10.6092 | -34.547 | -17.6953
10.6153 | -34.5474 | -17.7044
10.6214 | -34.5479 | -17.7134
10.6275 | -34.5483 | -17.7225
10.6336 | -34.5487 | -17.7318
10.6397 | -34.5492 | -17.7407
10.6458 | -34.5495 | -17.7496
10.6519 | -34.5499 | -17.7585
10.658 | -34.5504 | -17.7673
10.6641 | -34.5508 | -17.7762
10.6702 | -34.5513 | -17.7854
10.6764 | -34.5517 | -17.7943
10.6825 | -34.5521 | -17.8036
10.6886 | -34.5526 | -17.8129
10.6947 | -34.553 | -17.8217
10.7008 | -34.5534 | -17.8305
10.7069 | -34.5538 | -17.8396
10.713 | -34.5543 | -17.8485
10.7191 | -34.5547 | -17.8577
10.7252 | -34.5552 | -17.8668
10.7313 | -34.5556 | -17.8759
10.7374 | -34.556 | -17.8847
10.7435 | -34.5564 | -17.8939
10.7496 | -34.5569 | -17.903
10.7557 | -34.5573 | -17.9119
10.7618 | -34.5577 | -17.9212
10.7679 | -34.5582 | -17.9302
10.774 | -34.5586 | -17.939
10.7801 | -34.559 | -17.9485
10.7862 | -34.5594 | -17.9576
10.7923 | -34.5599 | -17.9666
10.7984 | -34.5603 | -17.9756
10.8045 | -34.5607 | -17.985
10.8106 | -34.5612 | -17.9939
10.8168 | -34.5616 | -18.0031
10.8229 | -34.562 | -18.0122
10.829 | -34.5625 | -18.0214
10.8351 | -34.5629 | -18.0306
10.8412 | -34.5632 | -18.0397
10.8473 | -34.5636 | -18.0486
10.8534 | -34.564 | -18.0578
10.8595 | -34.5645 | -18.0666
10.8656 | -34.5649 | -18.0757
10.8717 | -34.5654 | -18.0844
10.8778 | -34.5658 | -18.0932
10.8839 | -34.5662 | -18.1022
10.89 | -34.5666 | -18.111
10.8961 | -34.567 | -18.12
10.9022 | -34.5675 | -18.129
10.9083 | -34.5679 | -18.1383
10.9144 | -34.5684 | -18.1474
10.9205 | -34.5689 | -18.1564
10.9266 | -34.5693 | -18.1658
10.9327 | -34.5698 | -18.1748
10.9388 | -34.5702 | -18.1839
10.9449 | -34.5706 | -18.193
10.951 | -34.571 | -18.2019
10.9571 | -34.5714 | -18.211
10.9633 | -34.5719 | -18.2202
10.9694 | -34.5724 | -18.2294
10.9755 | -34.5728 | -18.2382
10.9816 | -34.5733 | -18.2472
10.9877 | -34.5737 | -18.2563
10.9938 | -34.5741 | -18.2654
10.9999 | -34.5745 | -18.2744
11.006 | -34.575 | -18.2834
11.0121 | -34.5754 | -18.2928
11.0182 | -34.5758 | -18.3015
11.0243 | -34.5763 | -18.3105
11.0304 | -34.5767 | -18.3195
11.0365 | -34.5771 | -18.3285
11.0426 | -34.5775 | -18.3373
11.0487 | -34.578 | -18.3462
11.0548 | -34.5784 | -18.3553
11.0609 | -34.5788 | -18.3645
11.067 | -34.5792 | -18.3736
11.0731 | -34.5797 | -18.3828
11.0792 | -34.5801 | -18.3918
11.0853 | -34.5805 | -18.4004
11.0914 | -34.581 | -18.4092
11.0975 | -34.5814 | -18.4185
11.1037 | -34.5819 | -18.4274
11.1098 | -34.5823 | -18.4364
11.1159 | -34.5827 | -18.4458
11.122 | -34.5832 | -18.4552
11.1281 | -34.5836 | -18.4639
11.1342 | -34.584 | -18.4725
11.1403 | -34.5844 | -18.4817
11.1464 | -34.5849 | -18.4905
11.1525 | -34.5853 | -18.4998
11.1586 | -34.5858 | -18.5092
11.1647 | -34.5862 | -18.5181
11.1708 | -34.5867 | -18.527
11.1769 | -34.5871 | -18.5356
11.183 | -34.5875 | -18.5445
11.1891 | -34.5879 | -18.5538
11.1952 | -34.5883 | -18.5631
11.2013 | -34.5887 | -18.572
11.2074 | -34.5892 | -18.5808
11.2135 | -34.5895 | -18.5896
11.2196 | -34.59 | -18.5985
11.2257 | -34.5903 | -18.6073
11.2318 | -34.5907 | -18.6166
11.2379 | -34.591 | -18.6253
11.244 | -34.5914 | -18.634
11.2502 | -34.5918 | -18.643
11.2563 | -34.5923 | -18.6518
11.2624 | -34.5927 | -18.661
11.2685 | -34.5931 | -18.6699
11.2746 | -34.5935 | -18.6787
11.2807 | -34.5939 | -18.6879
11.2868 | -34.5943 | -18.697
11.2929 | -34.5948 | -18.706
11.299 | -34.5952 | -18.715
11.3051 | -34.5956 | -18.7237
11.3112 | -34.596 | -18.7326
11.3173 | -34.5964 | -18.7418
11.3234 | -34.5968 | -18.751
11.3295 | -34.5972 | -18.7596
11.3356 | -34.5977 | -18.7683
11.3417 | -34.5981 | -18.7776
11.3478 | -34.5985 | -18.7864
11.3539 | -34.599 | -18.7952
11.36 | -34.5994 | -18.8041
11.3661 | -34.5998 | -18.813
11.3722 | -34.6002 | -18.8221
11.3783 | -34.6006 | -18.8308
11.3844 | -34.601 | -18.8398
11.3906 | -34.6014 | -18.849
11.3967 | -34.6018 | -18.8579
11.4028 | -34.6022 | -18.8669
11.4089 | -34.6027 | -18.8759
11.415 | -34.6031 | -18.8848
11.4211 | -34.6035 | -18.8934
11.4272 | -34.6039 | -18.9024
11.4333 | -34.6044 | -18.9115
11.4394 | -34.6048 | -18.9208
11.4455 | -34.6052 | -18.93
11.4516 | -34.6056 | -18.9387
11.4577 | -34.6061 | -18.9479
11.4638 | -34.6065 | -18.9565
11.4699 | -34.6069 | -18.9657
11.476 | -34.6073 | -18.9749
11.4821 | -34.6078 | -18.9836
11.4882 | -34.6082 | -18.9927
11.4943 | -34.6086 | -19.0014
11.5004 | -34.6091 | -19.0104
11.5065 | -34.6095 | -19.0194
11.5126 | -34.6099 | -19.0282
11.5187 | -34.6104 | -19.0371
11.5248 | -34.6108 | -19.0458
11.5309 | -34.6112 | -19.0547
11.5371 | -34.6116 | -19.0636
11.5432 | -34.612 | -19.0727
11.5493 | -34.6124 | -19.0816
11.5554 | -34.6128 | -19.0903
11.5615 | -34.6132 | -19.0993
11.5676 | -34.6137 | -19.1084
11.5737 | -34.6141 | -19.1171
11.5798 | -34.6145 | -19.126
11.5859 | -34.615 | -19.1348
11.592 | -34.6154 | -19.1439
11.5981 | -34.6158 | -19.1528
11.6042 | -34.6163 | -19.1617
11.6103 | -34.6167 | -19.1707
11.6164 | -34.6171 | -19.1796
11.6225 | -34.6175 | -19.1885
11.6286 | -34.6179 | -19.1976
11.6347 | -34.6184 | -19.2067
11.6408 | -34.6188 | -19.2156
11.6469 | -34.6193 | -19.2248
11.653 | -34.6197 | -19.2335
11.6591 | -34.6201 | -19.2425
11.6652 | -34.6205 | -19.2517
11.6713 | -34.6209 | -19.2607
11.6775 | -34.6214 | -19.2694
11.6836 | -34.6218 | -19.2785
11.6897 | -34.6222 | -19.2871
11.6958 | -34.6226 | -19.2963
11.7019 | -34.6231 | -19.305
11.708 | -34.6235 | -19.3142
11.7141 | -34.624 | -19.3231
11.7202 | -34.6244 | -19.3318
11.7263 | -34.6248 | -19.3408
11.7324 | -34.6253 | -19.3496
11.7385 | -34.6257 | -19.3587
11.7446 | -34.6261 | -19.368
11.7507 | -34.6265 | -19.3769
11.7568 | -34.627 | -19.386
11.7629 | -34.6274 | -19.3948
11.769 | -34.6278 | -19.4039
11.7751 | -34.6283 | -19.4126
11.7812 | -34.6287 | -19.4213
11.7873 | -34.6291 | -19.4305
11.7934 | -34.6296 | -19.4394
11.7995 | -34.63 | -19.448
11.8056 | -34.6304 | -19.4575
11.8117 | -34.6308 | -19.4665
11.8178 | -34.6312 | -19.4753
11.824 | -34.6317 | -19.4841
11.8301 | -34.6321 | -19.4927
11.8362 | -34.6325 | -19.5015
11.8423 | -34.6329 | -19.5105
11.8484 | -34.6334 | -19.5194
11.8545 | -34.6338 | -19.5285
11.8606 | -34.6342 | -19.5376
11.8667 | -34.6347 | -19.5464
11.8728 | -34.6351 | -19.5555
11.8789 | -34.6355 | -19.5649
11.885 | -34.6359 | -19.5736
11.8911 | -34.6364 | -19.5827
11.8972 | -34.6368 | -19.5913
11.9033 | -34.6372 | -19.5999
11.9094 | -34.6376 | -19.6086
11.9155 | -34.6379 | -19.6175
11.9216 | -34.6384 | -19.6262
11.9277 | -34.6389 | -19.6351
11.9338 | -34.6393 | -19.644
11.9399 | -34.6397 | -19.653
11.946 | -34.6401 | -19.662
11.9521 | -34.6405 | -19.6709
11.9582 | -34.641 | -19.6796
11.9644 | -34.6415 | -19.6885
11.9705 | -34.6418 | -19.6975
11.9766 | -34.6422 | -19.7064
11.9827 | -34.6426 | -19.7151
11.9888 | -34.643 | -19.7238
11.9949 | -34.6434 | -19.7329
12.001 | -34.6439 | -19.7418
12.0071 | -34.6443 | -19.7505
12.0132 | -34.6447 | -19.7593
12.0193 | -34.6452 | -19.7683
12.0254 | -34.6456 | -19.777
12.0315 | -34.646 | -19.7859
12.0376 | -34.6465 | -19.7947
12.0437 | -34.6469 | -19.8034
12.0498 | -34.6473 | -19.8124
12.0559 | -34.6477 | -19.8215
12.062 | -34.6481 | -19.8305
12.0681 | -34.6485 | -19.8392
12.0742 | -34.649 | -19.8484
12.0803 | -34.6494 | -19.857
12.0864 | -34.6499 | -19.8657
12.0925 | -34.6503 | -19.8746
12.0986 | -34.6508 | -19.8831
12.1047 | -34.6512 | -19.892
12.1109 | -34.6516 | -19.9011
12.117 | -34.6521 | -19.9099
12.1231 | -34.6525 | -19.9188
12.1292 | -34.6529 | -19.9279
12.1353 | -34.6534 | -19.9365
12.1414 | -34.6538 | -19.9454
12.1475 | -34.6542 | -19.9542
12.1536 | -34.6546 | -19.9631
12.1597 | -34.6551 | -19.972
12.1658 | -34.6555 | -19.9809
12.1719 | -34.656 | -19.9901
12.178 | -34.6564 | -19.9991
12.1841 | -34.6568 | -20.0079
12.1902 | -34.6573 | -20.0168
12.1963 | -34.6577 | -20.0255
12.2024 | -34.6582 | -20.0343
12.2085 | -34.6586 | -20.0432
12.2146 | -34.659 | -20.0519
12.2207 | -34.6594 | -20.0608
12.2268 | -34.6598 | -20.0696
12.2329 | -34.6603 | -20.0784
12.239 | -34.6606 | -20.0872
12.2451 | -34.6611 | -20.0957
12.2513 | -34.6616 | -20.1045
12.2574 | -34.662 | -20.1137
12.2635 | -34.6624 | -20.1221
12.2696 | -34.6628 | -20.1309
12.2757 | -34.6632 | -20.1395
12.2818 | -34.6636 | -20.1481
12.2879 | -34.664 | -20.1568
12.294 | -34.6644 | -20.1657
12.3001 | -34.6649 | -20.1747
12.3062 | -34.6653 | -20.1835
12.3123 | -34.6657 | -20.1925
12.3184 | -34.6661 | -20.2013
12.3245 | -34.6666 | -20.2105
12.3306 | -34.667 | -20.2192
12.3367 | -34.6674 | -20.2277
12.3428 | -34.6679 | -20.2363
12.3489 | -34.6683 | -20.2447
12.355 | -34.6687 | -20.2534
12.3611 | -34.6691 | -20.262
12.3672 | -34.6696 | -20.2708
12.3733 | -34.67 | -20.2799
12.3794 | -34.6704 | -20.2884
12.3855 | -34.6708 | -20.2973
12.3916 | -34.6712 | -20.3063
12.3978 | -34.6716 | -20.3152
12.4039 | -34.672 | -20.324
12.41 | -34.6724 | -20.3328
12.4161 | -34.6729 | -20.3416
12.4222 | -34.6733 | -20.3502
12.4283 | -34.6738 | -20.3591
12.4344 | -34.6742 | -20.3677
12.4405 | -34.6746 | -20.3765
12.4466 | -34.675 | -20.3852
12.4527 | -34.6755 | -20.3941
12.4588 | -34.6759 | -20.4029
12.4649 | -34.6763 | -20.4119
12.471 | -34.6767 | -20.421
12.4771 | -34.6772 | -20.4293
12.4832 | -34.6776 | -20.4381
12.4893 | -34.678 | -20.4467
12.4954 | -34.6785 | -20.4556
12.5015 | -34.679 | -20.4643
12.5076 | -34.6794 | -20.4729
12.5137 | -34.6797 | -20.482
12.5198 | -34.6801 | -20.4905
12.5259 | -34.6806 | -20.4991
12.532 | -34.6811 | -20.5081
12.5382 | -34.6815 | -20.5166
12.5443 | -34.682 | -20.5256
12.5504 | -34.6824 | -20.5343
12.5565 | -34.6829 | -20.5428
12.5626 | -34.6833 | -20.5517
12.5687 | -34.6837 | -20.5605
12.5748 | -34.6841 | -20.5691
12.5809 | -34.6846 | -20.5776
12.587 | -34.6849 | -20.5865
12.5931 | -34.6854 | -20.5953
12.5992 | -34.6858 | -20.6042
12.6053 | -34.6862 | -20.6129
12.6114 | -34.6867 | -20.6218
12.6175 | -34.6871 | -20.6305
12.6236 | -34.6876 | -20.639
12.6297 | -34.688 | -20.6475
12.6358 | -34.6885 | -20.6562
12.6419 | -34.689 | -20.6651
12.648 | -34.6893 | -20.674
12.6541 | -34.6898 | -20.6827
12.6602 | -34.6903 | -20.6912
12.6663 | -34.6906 | -20.6997
12.6724 | -34.6911 | -20.7086
12.6785 | -34.6915 | -20.7176
12.6847 | -34.6919 | -20.7266
12.6908 | -34.6924 | -20.7352
12.6969 | -34.6928 | -20.7438
12.703 | -34.6933 | -20.7526
12.7091 | -34.6937 | -20.7615
12.7152 | -34.6941 | -20.7704
12.7213 | -34.6945 | -20.7792
12.7274 | -34.695 | -20.788
12.7335 | -34.6954 | -20.7963
12.7396 | -34.6958 | -20.8048
12.7457 | -34.6963 | -20.8135
12.7518 | -34.6967 | -20.8223
12.7579 | -34.6971 | -20.8311
12.764 | -34.6975 | -20.8398
12.7701 | -34.6979 | -20.8481
12.7762 | -34.6983 | -20.8569
12.7823 | -34.6987 | -20.8655
12.7884 | -34.6991 | -20.8741
12.7945 | -34.6995 | -20.8822
12.8006 | -34.7 | -20.8902
12.8067 | -34.7004 | -20.8993
12.8128 | -34.7008 | -20.9077
12.8189 | -34.7011 | -20.9161
12.8251 | -34.7015 | -20.9246
12.8312 | -34.702 | -20.9331
12.8373 | -34.7024 | -20.9417
12.8434 | -34.7029 | -20.9501
12.8495 | -34.7032 | -20.9588
12.8556 | -34.7037 | -20.9675
12.8617 | -34.7041 | -20.9758
12.8678 | -34.7045 | -20.9843
12.8739 | -34.7049 | -20.9926
12.88 | -34.7054 | -21.0015
12.8861 | -34.7058 | -21.0103
12.8922 | -34.7062 | -21.0187
12.8983 | -34.7066 | -21.027
12.9044 | -34.7071 | -21.0357
12.9105 | -34.7076 | -21.0442
12.9166 | -34.708 | -21.0532
12.9227 | -34.7084 | -21.0623
12.9288 | -34.7089 | -21.071
12.9349 | -34.7093 | -21.0799
12.941 | -34.7098 | -21.0885
12.9471 | -34.7103 | -21.097
12.9532 | -34.7107 | -21.1057
12.9593 | -34.7112 | -21.1144
12.9654 | -34.7116 | -21.123
12.9716 | -34.712 | -21.1314
12.9777 | -34.7124 | -21.1394
12.9838 | -34.7128 | -21.1479
12.9899 | -34.7132 | -21.1567
12.996 | -34.7136 | -21.1655
13.0021 | -34.714 | -21.1741
13.0082 | -34.7144 | -21.1827
13.0143 | -34.7148 | -21.191
13.0204 | -34.7152 | -21.1994
13.0265 | -34.7156 | -21.2077
13.0326 | -34.7161 | -21.2166
13.0387 | -34.7165 | -21.2251
13.0448 | -34.7169 | -21.2338
13.0509 | -34.7174 | -21.2421
13.057 | -34.7178 | -21.2505
13.0631 | -34.7183 | -21.259
13.0692 | -34.7187 | -21.2675
13.0753 | -34.7191 | -21.2759
13.0814 | -34.7195 | -21.284
13.0875 | -34.72 | -21.2921
13.0936 | -34.7204 | -21.3004
13.0997 | -34.7208 | -21.3087
13.1058 | -34.7212 | -21.3166
13.112 | -34.7216 | -21.3251
13.1181 | -34.722 | -21.3336
13.1242 | -34.7224 | -21.3421
13.1303 | -34.7229 | -21.3509
13.1364 | -34.7233 | -21.3594
13.1425 | -34.7238 | -21.3679
13.1486 | -34.7242 | -21.3762
13.1547 | -34.7247 | -21.3842
13.1608 | -34.7251 | -21.3927
13.1669 | -34.7255 | -21.4013
13.173 | -34.7259 | -21.4096
13.1791 | -34.7264 | -21.4183
13.1852 | -34.7268 | -21.4271
13.1913 | -34.7272 | -21.4355
13.1974 | -34.7276 | -21.4439
13.2035 | -34.728 | -21.4524
13.2096 | -34.7285 | -21.4609
13.2157 | -34.7289 | -21.4695
13.2218 | -34.7294 | -21.4782
13.2279 | -34.7299 | -21.4865
13.234 | -34.7302 | -21.4948
13.2401 | -34.7307 | -21.5032
13.2462 | -34.7311 | -21.5118
13.2524 | -34.7315 | -21.52
13.2585 | -34.732 | -21.5284
13.2646 | -34.7324 | -21.537
13.2707 | -34.7328 | -21.5455
13.2768 | -34.7332 | -21.5541
13.2829 | -34.7337 | -21.5626
13.289 | -34.7341 | -21.5711
13.2951 | -34.7345 | -21.5793
13.3012 | -34.735 | -21.588
13.3073 | -34.7353 | -21.5965
13.3134 | -34.7358 | -21.6048
13.3195 | -34.7362 | -21.6132
13.3256 | -34.7367 | -21.6219
13.3317 | -34.7371 | -21.6301
13.3378 | -34.7375 | -21.6386
13.3439 | -34.7379 | -21.6468
13.35 | -34.7384 | -21.6551
13.3561 | -34.7388 | -21.6637
13.3622 | -34.7392 | -21.6722
13.3683 | -34.7397 | -21.6804
13.3744 | -34.7401 | -21.6889
13.3805 | -34.7405 | -21.6974
13.3866 | -34.7409 | -21.7058
13.3927 | -34.7413 | -21.7142
13.3989 | -34.7418 | -21.7223
13.405 | -34.7423 | -21.7307
13.4111 | -34.7428 | -21.7393
13.4172 | -34.7432 | -21.7474
13.4233 | -34.7436 | -21.7559
13.4294 | -34.744 | -21.7639
13.4355 | -34.7445 | -21.773
13.4416 | -34.7449 | -21.7815
13.4477 | -34.7452 | -21.7898
13.4538 | -34.7456 | -21.7982
13.4599 | -34.7461 | -21.8065
13.466 | -34.7465 | -21.8152
13.4721 | -34.747 | -21.8237
13.4782 | -34.7475 | -21.832
13.4843 | -34.7479 | -21.8402
13.4904 | -34.7484 | -21.8482
13.4965 | -34.7488 | -21.8568
13.5026 | -34.7492 | -21.8652
13.5087 | -34.7497 | -21.8733
13.5148 | -34.7501 | -21.8815
13.5209 | -34.7506 | -21.8897
13.527 | -34.751 | -21.8985
13.5331 | -34.7514 | -21.9066
13.5393 | -34.7518 | -21.9148
13.5454 | -34.7522 | -21.9234
13.5515 | -34.7527 | -21.9316
13.5576 | -34.7531 | -21.94
13.5637 | -34.7535 | -21.9482
13.5698 | -34.754 | -21.9567
13.5759 | -34.7543 | -21.9648
13.582 | -34.7547 | -21.9734
13.5881 | -34.7552 | -21.9818
13.5942 | -34.7556 | -21.9901
13.6003 | -34.7561 | -21.9984
13.6064 | -34.7565 | -22.0069
13.6125 | -34.757 | -22.0153
13.6186 | -34.7574 | -22.0238
13.6247 | -34.7578 | -22.032
13.6308 | -34.7582 | -22.04
13.6369 | -34.7587 | -22.0484
13.643 | -34.7591 | -22.0565
13.6491 | -34.7596 | -22.0649
13.6552 | -34.76 | -22.0731
13.6613 | -34.7604 | -22.0816
13.6674 | -34.7609 | -22.0899
13.6735 | -34.7613 | -22.0984
13.6796 | -34.7618 | -22.1065
13.6858 | -34.7621 | -22.1149
13.6919 | -34.7626 | -22.1233
13.698 | -34.763 | -22.1317
13.7041 | -34.7634 | -22.1401
13.7102 | -34.7639 | -22.1486
13.7163 | -34.7643 | -22.157
13.7224 | -34.7648 | -22.1652
13.7285 | -34.7652 | -22.1736
13.7346 | -34.7657 | -22.1818
13.7407 | -34.7661 | -22.1901
13.7468 | -34.7665 | -22.1984
13.7529 | -34.7669 | -22.2066
13.759 | -34.7674 | -22.2151
13.7651 | -34.7678 | -22.2235
13.7712 | -34.7682 | -22.2319
13.7773 | -34.7687 | -22.2402
13.7834 | -34.7691 | -22.2484
13.7895 | -34.7696 | -22.2563
13.7956 | -34.77 | -22.2648
13.8017 | -34.7705 | -22.2733
13.8078 | -34.771 | -22.2817
13.8139 | -34.7715 | -22.29
13.82 | -34.7719 | -22.2982
13.8262 | -34.7724 | -22.3067
13.8323 | -34.7728 | -22.3151
13.8384 | -34.7733 | -22.3235
13.8445 | -34.7737 | -22.3318
13.8506 | -34.7741 | -22.3404
13.8567 | -34.7746 | -22.3488
13.8628 | -34.775 | -22.3572
13.8689 | -34.7754 | -22.3654
13.875 | -34.7759 | -22.3738
13.8811 | -34.7764 | -22.3821
13.8872 | -34.7769 | -22.3905
13.8933 | -34.7774 | -22.3987
13.8994 | -34.7778 | -22.4068
13.9055 | -34.7783 | -22.4148
13.9116 | -34.7787 | -22.4228
13.9177 | -34.7792 | -22.4308
13.9238 | -34.7797 | -22.4392
13.9299 | -34.7802 | -22.4475
13.936 | -34.7806 | -22.4557
13.9421 | -34.7811 | -22.4641
13.9482 | -34.7816 | -22.4726
13.9543 | -34.7821 | -22.4805
13.9604 | -34.7825 | -22.4889
13.9665 | -34.783 | -22.4971
13.9727 | -34.7834 | -22.505
13.9788 | -34.7838 | -22.5135
13.9849 | -34.7842 | -22.5217
13.991 | -34.7846 | -22.5303
13.9971 | -34.7851 | -22.539
14.0032 | -34.7855 | -22.5476
14.0093 | -34.7859 | -22.5557
14.0154 | -34.7864 | -22.5644
14.0215 | -34.7868 | -22.5725
14.0276 | -34.7872 | -22.5806
14.0337 | -34.7877 | -22.5889
14.0398 | -34.7881 | -22.5971
14.0459 | -34.7885 | -22.6055
14.052 | -34.7889 | -22.6137
14.0581 | -34.7894 | -22.6217
14.0642 | -34.7899 | -22.6302
14.0703 | -34.7904 | -22.6383
14.0764 | -34.7908 | -22.6471
14.0825 | -34.7913 | -22.6554
14.0886 | -34.7916 | -22.6631
14.0947 | -34.792 | -22.6709
14.1008 | -34.7925 | -22.6796
14.1069 | -34.793 | -22.6877
14.1131 | -34.7934 | -22.6961
14.1192 | -34.7938 | -22.7042
14.1253 | -34.7943 | -22.7124
14.1314 | -34.7948 | -22.7205
14.1375 | -34.7953 | -22.7289
14.1436 | -34.7957 | -22.7371
14.1497 | -34.7962 | -22.7452
14.1558 | -34.7967 | -22.7537
14.1619 | -34.7971 | -22.762
14.168 | -34.7976 | -22.7701
14.1741 | -34.798 | -22.7782
14.1802 | -34.7985 | -22.7864
14.1863 | -34.7989 | -22.7946
14.1924 | -34.7993 | -22.8028
14.1985 | -34.7998 | -22.8108
14.2046 | -34.8003 | -22.8188
14.2107 | -34.8007 | -22.8267
14.2168 | -34.8012 | -22.8353
14.2229 | -34.8017 | -22.8435
14.229 | -34.8021 | -22.8516
14.2351 | -34.8026 | -22.8596
14.2412 | -34.803 | -22.8675
14.2473 | -34.8035 | -22.8758
14.2534 | -34.804 | -22.884
14.2596 | -34.8045 | -22.8923
14.2657 | -34.8049 | -22.9007
14.2718 | -34.8054 | -22.909
14.2779 | -34.8058 | -22.9168
14.284 | -34.8063 | -22.9247
14.2901 | -34.8068 | -22.9328
14.2962 | -34.8073 | -22.9407
14.3023 | -34.8077 | -22.9489
14.3084 | -34.8082 | -22.9574
14.3145 | -34.8087 | -22.9654
14.3206 | -34.8091 | -22.9735
14.3267 | -34.8096 | -22.9816
14.3328 | -34.81 | -22.9896
14.3389 | -34.8105 | -22.9976
14.345 | -34.8109 | -23.0059
14.3511 | -34.8114 | -23.0137
14.3572 | -34.8118 | -23.0219
14.3633 | -34.8123 | -23.0299
14.3694 | -34.8127 | -23.038
14.3755 | -34.8131 | -23.0458
14.3816 | -34.8136 | -23.0537
14.3877 | -34.8141 | -23.0615
14.3938 | -34.8146 | -23.0697
14.4 | -34.815 | -23.0777
14.4061 | -34.8154 | -23.0858
14.4122 | -34.8158 | -23.0942
14.4183 | -34.8163 | -23.1025
14.4244 | -34.8167 | -23.1102
14.4305 | -34.8172 | -23.1186
14.4366 | -34.8176 | -23.1268
14.4427 | -34.8181 | -23.1349
14.4488 | -34.8185 | -23.1426
14.4549 | -34.819 | -23.1507
14.461 | -34.8195 | -23.1591
14.4671 | -34.8199 | -23.1672
14.4732 | -34.8204 | -23.1749
14.4793 | -34.8209 | -23.1836
14.4854 | -34.8214 | -23.1913
14.4915 | -34.8218 | -23.1993
14.4976 | -34.8223 | -23.2076
14.5037 | -34.8228 | -23.2159
14.5098 | -34.8232 | -23.2239
14.5159 | -34.8237 | -23.2322
14.522 | -34.8242 | -23.2406
14.5281 | -34.8247 | -23.2484
14.5342 | -34.8251 | -23.2565
14.5403 | -34.8256 | -23.2649
14.5465 | -34.8261 | -23.2733
14.5526 | -34.8265 | -23.2816
14.5587 | -34.8269 | -23.2896
14.5648 | -34.8274 | -23.2975
14.5709 | -34.8279 | -23.3059
14.577 | -34.8284 | -23.3142
14.5831 | -34.8288 | -23.3224
14.5892 | -34.8294 | -23.3301
14.5953 | -34.8298 | -23.3378
14.6014 | -34.8303 | -23.3459
14.6075 | -34.8308 | -23.3538
14.6136 | -34.8313 | -23.3618
14.6197 | -34.8317 | -23.3699
14.6258 | -34.8322 | -23.378
14.6319 | -34.8327 | -23.3858
14.638 | -34.8332 | -23.3939
14.6441 | -34.8337 | -23.4015
14.6502 | -34.8341 | -23.4091
14.6563 | -34.8346 | -23.417
14.6624 | -34.8351 | -23.425
14.6685 | -34.8355 | -23.4326
14.6746 | -34.836 | -23.4406
14.6807 | -34.8364 | -23.4488
14.6869 | -34.8369 | -23.4567
14.693 | -34.8373 | -23.4646
14.6991 | -34.8378 | -23.4726
14.7052 | -34.8383 | -23.4804
14.7113 | -34.8387 | -23.4884
14.7174 | -34.8392 | -23.4964
14.7235 | -34.8396 | -23.5043
14.7296 | -34.8401 | -23.5123
14.7357 | -34.8406 | -23.52
14.7418 | -34.841 | -23.5277
14.7479 | -34.8415 | -23.5358
14.754 | -34.842 | -23.5437
14.7601 | -34.8425 | -23.552
14.7662 | -34.8429 | -23.5601
14.7723 | -34.8434 | -23.568
14.7784 | -34.8439 | -23.5761
14.7845 | -34.8444 | -23.5841
14.7906 | -34.8448 | -23.5919
14.7967 | -34.8453 | -23.6
14.8028 | -34.8457 | -23.6079
14.8089 | -34.8462 | -23.6156
14.815 | -34.8466 | -23.6233
14.8211 | -34.847 | -23.6311
14.8272 | -34.8475 | -23.6392
14.8334 | -34.848 | -23.6471
14.8395 | -34.8484 | -23.6549
14.8456 | -34.8489 | -23.6628
14.8517 | -34.8494 | -23.6708
14.8578 | -34.8499 | -23.6786
14.8639 | -34.8503 | -23.6867
14.87 | -34.8508 | -23.6945
14.8761 | -34.8513 | -23.702
14.8822 | -34.8517 | -23.7099
14.8883 | -34.8522 | -23.7177
14.8944 | -34.8527 | -23.7258
14.9005 | -34.8532 | -23.7335
14.9066 | -34.8536 | -23.7414
14.9127 | -34.8541 | -23.7491
14.9188 | -34.8546 | -23.7574
14.9249 | -34.855 | -23.7653
14.931 | -34.8555 | -23.773
14.9371 | -34.8559 | -23.7808
14.9432 | -34.8564 | -23.7887
14.9493 | -34.8569 | -23.7966
14.9554 | -34.8574 | -23.8044
14.9615 | -34.8578 | -23.8124
14.9676 | -34.8584 | -23.8204
14.9738 | -34.8588 | -23.8282
14.9799 | -34.8593 | -23.8362
14.986 | -34.8598 | -23.8438
14.9921 | -34.8603 | -23.8516
14.9982 | -34.8607 | -23.8598
15.0043 | -34.8612 | -23.8678
15.0104 | -34.8616 | -23.8755
15.0165 | -34.8621 | -23.8833
15.0226 | -34.8626 | -23.8908
15.0287 | -34.863 | -23.8991
15.0348 | -34.8635 | -23.907
15.0409 | -34.864 | -23.9151
15.047 | -34.8645 | -23.923
15.0531 | -34.865 | -23.931
15.0592 | -34.8655 | -23.9389
15.0653 | -34.866 | -23.9467
15.0714 | -34.8665 | -23.9544
15.0775 | -34.8669 | -23.962
15.0836 | -34.8674 | -23.9699
15.0897 | -34.8679 | -23.9776
15.0958 | -34.8683 | -23.9853
15.1019 | -34.8687 | -23.9932
15.108 | -34.8692 | -24.0012
15.1141 | -34.8696 | -24.0092
15.1203 | -34.8702 | -24.0169
15.1264 | -34.8707 | -24.0248
15.1325 | -34.8712 | -24.0323
15.1386 | -34.8716 | -24.0403
15.1447 | -34.8721 | -24.0481
15.1508 | -34.8726 | -24.0561
15.1569 | -34.8731 | -24.0641
15.163 | -34.8736 | -24.0716
15.1691 | -34.8741 | -24.0794
15.1752 | -34.8746 | -24.0873
15.1813 | -34.8751 | -24.0952
15.1874 | -34.8756 | -24.1032
15.1935 | -34.8761 | -24.1106
15.1996 | -34.8766 | -24.1188
15.2057 | -34.8771 | -24.1265
15.2118 | -34.8775 | -24.1344
15.2179 | -34.878 | -24.1419
15.224 | -34.8785 | -24.1499
15.2301 | -34.879 | -24.1575
15.2362 | -34.8795 | -24.1653
15.2423 | -34.88 | -24.1733
15.2484 | -34.8804 | -24.1809
15.2545 | -34.8809 | -24.1885
15.2607 | -34.8814 | -24.1964
15.2668 | -34.8819 | -24.204
15.2729 | -34.8823 | -24.2117
15.279 | -34.8828 | -24.2195
15.2851 | -34.8833 | -24.2274
15.2912 | -34.8838 | -24.2349
15.2973 | -34.8842 | -24.2425
15.3034 | -34.8848 | -24.2499
15.3095 | -34.8852 | -24.2575
15.3156 | -34.8856 | -24.2655
15.3217 | -34.8861 | -24.2729
15.3278 | -34.8866 | -24.2805
15.3339 | -34.8871 | -24.288
15.34 | -34.8876 | -24.2957
15.3461 | -34.888 | -24.3037
15.3522 | -34.8885 | -24.3114
15.3583 | -34.889 | -24.319
15.3644 | -34.8895 | -24.327
15.3705 | -34.8899 | -24.3345
15.3766 | -34.8904 | -24.3425
15.3827 | -34.8908 | -24.35
15.3888 | -34.8913 | -24.3574
15.3949 | -34.8918 | -24.3649
15.401 | -34.8923 | -24.3728
15.4072 | -34.8928 | -24.3802
15.4133 | -34.8933 | -24.3881
15.4194 | -34.8938 | -24.3957
15.4255 | -34.8943 | -24.4034
15.4316 | -34.8948 | -24.4112
15.4377 | -34.8953 | -24.4186
15.4438 | -34.8958 | -24.4258
15.4499 | -34.8962 | -24.4335
15.456 | -34.8966 | -24.4413
15.4621 | -34.8971 | -24.449
15.4682 | -34.8977 | -24.4567
15.4743 | -34.8982 | -24.4642
15.4804 | -34.8987 | -24.472
15.4865 | -34.8992 | -24.4798
15.4926 | -34.8996 | -24.4875
15.4987 | -34.9001 | -24.4952
15.5048 | -34.9006 | -24.5027
15.5109 | -34.9011 | -24.51
15.517 | -34.9016 | -24.5177
15.5231 | -34.9021 | -24.5253
15.5292 | -34.9025 | -24.5329
15.5353 | -34.903 | -24.5403
15.5414 | -34.9035 | -24.5483
15.5476 | -34.904 | -24.556
15.5537 | -34.9045 | -24.5635
15.5598 | -34.905 | -24.5708
15.5659 | -34.9055 | -24.5787
15.572 | -34.9059 | -24.586
15.5781 | -34.9064 | -24.5939
15.5842 | -34.9069 | -24.6012
15.5903 | -34.9073 | -24.6087
15.5964 | -34.9078 | -24.6162
15.6025 | -34.9083 | -24.6238
15.6086 | -34.9088 | -24.6315
15.6147 | -34.9093 | -24.639
15.6208 | -34.9098 | -24.6463
15.6269 | -34.9103 | -24.654
15.633 | -34.9108 | -24.6615
15.6391 | -34.9112 | -24.6691
15.6452 | -34.9117 | -24.6766
15.6513 | -34.9122 | -24.6845
15.6574 | -34.9127 | -24.6916
15.6635 | -34.9131 | -24.6994
15.6696 | -34.9136 | -24.7068
15.6757 | -34.9141 | -24.7146
15.6818 | -34.9146 | -24.7225
15.688 | -34.9152 | -24.7301
15.6941 | -34.9156 | -24.7379
15.7002 | -34.9161 | -24.7455
15.7063 | -34.9166 | -24.7534
15.7124 | -34.9171 | -24.761
15.7185 | -34.9176 | -24.7685
15.7246 | -34.918 | -24.7757
15.7307 | -34.9185 | -24.7835
15.7368 | -34.919 | -24.7914
15.7429 | -34.9195 | -24.7989
15.749 | -34.9199 | -24.8064
15.7551 | -34.9204 | -24.8139
15.7612 | -34.9209 | -24.8207
15.7673 | -34.9213 | -24.8282
15.7734 | -34.9217 | -24.8355
15.7795 | -34.9222 | -24.8435
15.7856 | -34.9227 | -24.851
15.7917 | -34.9232 | -24.8585
15.7978 | -34.9237 | -24.8659
15.8039 | -34.9241 | -24.8734
15.81 | -34.9246 | -24.8808
15.8161 | -34.9251 | -24.8887
15.8222 | -34.9256 | -24.8964
15.8283 | -34.926 | -24.9038
15.8345 | -34.9265 | -24.9114
15.8406 | -34.927 | -24.9191
15.8467 | -34.9275 | -24.9265
15.8528 | -34.928 | -24.934
15.8589 | -34.9285 | -24.9415
15.865 | -34.9289 | -24.949
15.8711 | -34.9294 | -24.9568
15.8772 | -34.9299 | -24.9645
15.8833 | -34.9304 | -24.9723
15.8894 | -34.9309 | -24.9797
15.8955 | -34.9314 | -24.9875
15.9016 | -34.9319 | -24.9952
15.9077 | -34.9324 | -25.0027
15.9138 | -34.9329 | -25.0103
15.9199 | -34.9334 | -25.018
15.926 | -34.9339 | -25.0259
15.9321 | -34.9344 | -25.0335
15.9382 | -34.9349 | -25.0412
15.9443 | -34.9354 | -25.049
15.9504 | -34.9359 | -25.0564
15.9565 | -34.9364 | -25.0641
15.9626 | -34.9369 | -25.0716
15.9687 | -34.9374 | -25.079
15.9749 | -34.9379 | -25.0868
15.981 | -34.9384 | -25.0944
15.9871 | -34.9389 | -25.1018
15.9932 | -34.9394 | -25.1096
15.9993 | -34.9399 | -25.117
16.0054 | -34.9403 | -25.1245
16.0115 | -34.9408 | -25.1318
16.0176 | -34.9413 | -25.1394
16.0237 | -34.9418 | -25.1472
16.0298 | -34.9423 | -25.1545
16.0359 | -34.9428 | -25.162
16.042 | -34.9433 | -25.1693
16.0481 | -34.9438 | -25.1768
16.0542 | -34.9442 | -25.1846
16.0603 | -34.9448 | -25.1922
16.0664 | -34.9452 | -25.1998
16.0725 | -34.9458 | -25.2072
16.0786 | -34.9463 | -25.2151
16.0847 | -34.9467 | -25.2222
16.0908 | -34.9472 | -25.2295
16.0969 | -34.9477 | -25.2367
16.103 | -34.9482 | -25.2445
16.1091 | -34.9486 | -25.2518
16.1152 | -34.9491 | -25.2595
16.1214 | -34.9496 | -25.2671
16.1275 | -34.9501 | -25.2744
16.1336 | -34.9506 | -25.282
16.1397 | -34.9511 | -25.2894
16.1458 | -34.9516 | -25.2967
16.1519 | -34.9521 | -25.3043
16.158 | -34.9526 | -25.312
16.1641 | -34.9531 | -25.3193
16.1702 | -34.9535 | -25.3269
16.1763 | -34.954 | -25.3344
16.1824 | -34.9545 | -25.3416
16.1885 | -34.955 | -25.3493
16.1946 | -34.9555 | -25.3568
16.2007 | -34.956 | -25.3647
16.2068 | -34.9565 | -25.3719
16.2129 | -34.9569 | -25.3795
16.219 | -34.9573 | -25.387
16.2251 | -34.9578 | -25.3948
16.2312 | -34.9582 | -25.4024
16.2373 | -34.9587 | -25.4099
16.2434 | -34.9591 | -25.4175
16.2495 | -34.9597 | -25.4247
16.2556 | -34.9602 | -25.4321
16.2618 | -34.9607 | -25.4394
16.2679 | -34.9612 | -25.4469
16.274 | -34.9617 | -25.4547
16.2801 | -34.9622 | -25.4621
16.2862 | -34.9627 | -25.4693
16.2923 | -34.9631 | -25.4769
16.2984 | -34.9636 | -25.4846
16.3045 | -34.9642 | -25.4922
16.3106 | -34.9647 | -25.4993
16.3167 | -34.9651 | -25.5069
16.3228 | -34.9656 | -25.5145
16.3289 | -34.9661 | -25.5221
16.335 | -34.9665 | -25.5297
16.3411 | -34.967 | -25.5371
16.3472 | -34.9675 | -25.5445
16.3533 | -34.968 | -25.5521
16.3594 | -34.9684 | -25.5595
16.3655 | -34.9689 | -25.5672
16.3716 | -34.9694 | -25.5746
16.3777 | -34.9699 | -25.5821
16.3838 | -34.9704 | -25.5896
16.3899 | -34.9708 | -25.597
16.396 | -34.9713 | -25.6047
16.4021 | -34.9718 | -25.6123
16.4083 | -34.9723 | -25.6198
16.4144 | -34.9728 | -25.6271
16.4205 | -34.9733 | -25.6346
16.4266 | -34.9738 | -25.6423
16.4327 | -34.9743 | -25.6496
16.4388 | -34.9748 | -25.6572
16.4449 | -34.9753 | -25.6646
16.451 | -34.9758 | -25.6722
16.4571 | -34.9763 | -25.6798
16.4632 | -34.9768 | -25.6871
16.4693 | -34.9772 | -25.6947
16.4754 | -34.9777 | -25.702
16.4815 | -34.9782 | -25.7097
16.4876 | -34.9787 | -25.7173
16.4937 | -34.9792 | -25.7249
16.4998 | -34.9797 | -25.7322
16.5059 | -34.9802 | -25.7396
16.512 | -34.9807 | -25.7473
16.5181 | -34.9812 | -25.7548
16.5242 | -34.9816 | -25.7622
16.5303 | -34.9821 | -25.7695
16.5364 | -34.9826 | -25.777
16.5425 | -34.9831 | -25.7846
16.5487 | -34.9836 | -25.7921
16.5548 | -34.984 | -25.7995
16.5609 | -34.9845 | -25.8071
16.567 | -34.985 | -25.8145
16.5731 | -34.9854 | -25.8218
16.5792 | -34.9859 | -25.8292
16.5853 | -34.9864 | -25.8365
16.5914 | -34.9869 | -25.8441
16.5975 | -34.9874 | -25.8517
16.6036 | -34.9879 | -25.8592
16.6097 | -34.9884 | -25.8666
16.6158 | -34.9889 | -25.8743
16.6219 | -34.9894 | -25.8816
16.628 | -34.9898 | -25.8893
16.6341 | -34.9904 | -25.8967
16.6402 | -34.9909 | -25.904
16.6463 | -34.9914 | -25.9116
16.6524 | -34.9918 | -25.9191
16.6585 | -34.9923 | -25.9265
16.6646 | -34.9928 | -25.934
16.6707 | -34.9933 | -25.9413
16.6768 | -34.9938 | -25.9488
16.6829 | -34.9943 | -25.9563
16.689 | -34.9947 | -25.9638
16.6952 | -34.9952 | -25.9714
16.7013 | -34.9957 | -25.979
16.7074 | -34.9962 | -25.9864
16.7135 | -34.9967 | -25.9938
16.7196 | -34.9971 | -26.0013
16.7257 | -34.9976 | -26.0085
16.7318 | -34.998 | -26.0161
16.7379 | -34.9985 | -26.0237
16.744 | -34.999 | -26.0309
16.7501 | -34.9995 | -26.0383
16.7562 | -35 | -26.0457
16.7623 | -35.0005 | -26.0532
16.7684 | -35.001 | -26.0608
16.7745 | -35.0015 | -26.0678
16.7806 | -35.0019 | -26.0754
16.7867 | -35.0024 | -26.0826
16.7928 | -35.0028 | -26.0901
16.7989 | -35.0033 | -26.0977
16.805 | -35.0037 | -26.1052
16.8111 | -35.0043 | -26.1125
16.8172 | -35.0047 | -26.1203
16.8233 | -35.0052 | -26.1278
16.8294 | -35.0057 | -26.1353
16.8356 | -35.0062 | -26.1429
16.8417 | -35.0067 | -26.1504
16.8478 | -35.0071 | -26.1577
16.8539 | -35.0077 | -26.1652
16.86 | -35.0081 | -26.1723
16.8661 | -35.0086 | -26.1798
16.8722 | -35.0091 | -26.1871
16.8783 | -35.0096 | -26.194
16.8844 | -35.0101 | -26.2015
16.8905 | -35.0106 | -26.209
16.8966 | -35.011 | -26.2167
16.9027 | -35.0115 | -26.2238
16.9088 | -35.012 | -26.2316
16.9149 | -35.0125 | -26.2393
16.921 | -35.013 | -26.2467
16.9271 | -35.0135 | -26.2544
16.9332 | -35.014 | -26.2618
16.9393 | -35.0145 | -26.2696
16.9454 | -35.0149 | -26.277
16.9515 | -35.0154 | -26.2845
16.9576 | -35.0159 | -26.2919
16.9637 | -35.0164 | -26.2994
16.9698 | -35.0169 | -26.3067
16.9759 | -35.0174 | -26.314
16.9821 | -35.0178 | -26.3216
16.9882 | -35.0183 | -26.3291
16.9943 | -35.0188 | -26.3369
17.0004 | -35.0192 | -26.3443
17.0065 | -35.0197 | -26.3517
17.0126 | -35.0201 | -26.359
17.0187 | -35.0207 | -26.3665
17.0248 | -35.0211 | -26.3733
17.0309 | -35.0216 | -26.3807
17.037 | -35.0221 | -26.3883
17.0431 | -35.0226 | -26.3953
17.0492 | -35.023 | -26.4028
17.0553 | -35.0235 | -26.4105
17.0614 | -35.024 | -26.4179
17.0675 | -35.0244 | -26.4256
17.0736 | -35.0249 | -26.433
17.0797 | -35.0254 | -26.4402
17.0858 | -35.0259 | -26.4477
17.0919 | -35.0263 | -26.4553
17.098 | -35.0268 | -26.4624
17.1041 | -35.0273 | -26.4697
17.1102 | -35.0277 | -26.4773
17.1163 | -35.0282 | -26.485
17.1225 | -35.0287 | -26.4925
17.1286 | -35.0292 | -26.4996
17.1347 | -35.0296 | -26.5069
17.1408 | -35.0301 | -26.514
17.1469 | -35.0306 | -26.5214
17.153 | -35.0311 | -26.5289
17.1591 | -35.0316 | -26.5361
17.1652 | -35.032 | -26.5434
17.1713 | -35.0325 | -26.5509
17.1774 | -35.033 | -26.5585
17.1835 | -35.0334 | -26.5661
17.1896 | -35.0339 | -26.5735
17.1957 | -35.0344 | -26.5809
17.2018 | -35.0348 | -26.5885
17.2079 | -35.0353 | -26.5958
17.214 | -35.0357 | -26.6031
17.2201 | -35.0362 | -26.6105
17.2262 | -35.0367 | -26.6179
17.2323 | -35.0371 | -26.6253
17.2384 | -35.0377 | -26.6323
17.2445 | -35.0381 | -26.6398
17.2506 | -35.0386 | -26.6473
17.2567 | -35.039 | -26.6547
17.2628 | -35.0395 | -26.6623
17.269 | -35.0399 | -26.6695
17.2751 | -35.0404 | -26.6769
17.2812 | -35.0409 | -26.6842
17.2873 | -35.0414 | -26.6917
17.2934 | -35.0418 | -26.6989
17.2995 | -35.0422 | -26.706
17.3056 | -35.0427 | -26.7132
17.3117 | -35.0432 | -26.7204
17.3178 | -35.0436 | -26.7279
17.3239 | -35.0441 | -26.7351
17.33 | -35.0445 | -26.7424
17.3361 | -35.045 | -26.7497
17.3422 | -35.0455 | -26.757
17.3483 | -35.0459 | -26.7645
17.3544 | -35.0464 | -26.7717
17.3605 | -35.0469 | -26.7791
17.3666 | -35.0474 | -26.7867
17.3727 | -35.0479 | -26.794
17.3788 | -35.0483 | -26.8012
17.3849 | -35.0488 | -26.8087
17.391 | -35.0493 | -26.8159
17.3971 | -35.0499 | -26.8235
17.4032 | -35.0503 | -26.8312
17.4094 | -35.0508 | -26.8384
17.4155 | -35.0512 | -26.8458
17.4216 | -35.0517 | -26.8532
17.4277 | -35.0522 | -26.8606
17.4338 | -35.0527 | -26.868
17.4399 | -35.0531 | -26.8753
17.446 | -35.0536 | -26.8825
17.4521 | -35.0541 | -26.8895
17.4582 | -35.0545 | -26.8968
17.4643 | -35.055 | -26.9044
17.4704 | -35.0554 | -26.9116
17.4765 | -35.0559 | -26.919
17.4826 | -35.0564 | -26.9264
17.4887 | -35.0568 | -26.934
17.4948 | -35.0573 | -26.9414
17.5009 | -35.0577 | -26.9486
17.507 | -35.0582 | -26.956
17.5131 | -35.0587 | -26.9632
17.5192 | -35.0592 | -26.9704
17.5253 | -35.0596 | -26.9777
17.5314 | -35.0601 | -26.9851
17.5375 | -35.0606 | -26.9927
17.5436 | -35.061 | -27
17.5497 | -35.0615 | -27.0073
17.5559 | -35.0619 | -27.0147
17.562 | -35.0624 | -27.0219
17.5681 | -35.0628 | -27.0291
17.5742 | -35.0633 | -27.0365
17.5803 | -35.0638 | -27.0436
17.5864 | -35.0643 | -27.0505
17.5925 | -35.0648 | -27.0581
17.5986 | -35.0652 | -27.0651
17.6047 | -35.0657 | -27.0725
17.6108 | -35.0662 | -27.0796
17.6169 | -35.0667 | -27.0868
17.623 | -35.0671 | -27.0939
17.6291 | -35.0676 | -27.1014
17.6352 | -35.0681 | -27.1085
17.6413 | -35.0685 | -27.116
17.6474 | -35.069 | -27.1234
17.6535 | -35.0694 | -27.1308
17.6596 | -35.0699 | -27.138
17.6657 | -35.0704 | -27.1451
17.6718 | -35.0709 | -27.1527
17.6779 | -35.0713 | -27.1601
17.684 | -35.0718 | -27.1673
17.6901 | -35.0723 | -27.1744
17.6963 | -35.0728 | -27.1815
17.7024 | -35.0732 | -27.1888
17.7085 | -35.0737 | -27.1962
17.7146 | -35.0741 | -27.2036
17.7207 | -35.0746 | -27.2108
17.7268 | -35.0751 | -27.2177
17.7329 | -35.0755 | -27.2253
17.739 | -35.076 | -27.2327
17.7451 | -35.0765 | -27.24
17.7512 | -35.0769 | -27.2474
17.7573 | -35.0774 | -27.2547
17.7634 | -35.0779 | -27.2616
17.7695 | -35.0783 | -27.269
17.7756 | -35.0788 | -27.2763
17.7817 | -35.0793 | -27.2833
17.7878 | -35.0798 | -27.2906
17.7939 | -35.0802 | -27.2976
17.8 | -35.0807 | -27.3047
17.8061 | -35.0812 | -27.312
17.8122 | -35.0816 | -27.3192
17.8183 | -35.0821 | -27.3266
17.8244 | -35.0825 | -27.3339
17.8305 | -35.083 | -27.3413
17.8366 | -35.0835 | -27.3484
17.8428 | -35.084 | -27.3558
17.8489 | -35.0844 | -27.363
17.855 | -35.0849 | -27.3701
17.8611 | -35.0854 | -27.3773
17.8672 | -35.0858 | -27.3843
17.8733 | -35.0863 | -27.3917
17.8794 | -35.0868 | -27.3991
17.8855 | -35.0873 | -27.4063
17.8916 | -35.0877 | -27.4134
17.8977 | -35.0882 | -27.4205
17.9038 | -35.0887 | -27.4276
17.9099 | -35.0891 | -27.4349
17.916 | -35.0896 | -27.4419
17.9221 | -35.09 | -27.4491
17.9282 | -35.0905 | -27.4563
17.9343 | -35.091 | -27.4636
17.9404 | -35.0915 | -27.4707
17.9465 | -35.0919 | -27.4781
17.9526 | -35.0924 | -27.4854
17.9587 | -35.0929 | -27.4926
17.9648 | -35.0933 | -27.4998
17.9709 | -35.0938 | -27.5071
17.977 | -35.0943 | -27.5142
17.9832 | -35.0948 | -27.5215
17.9893 | -35.0952 | -27.5287
17.9954 | -35.0957 | -27.5361
18.0015 | -35.0963 | -27.5434
18.0076 | -35.0967 | -27.5501
18.0137 | -35.0972 | -27.5578
18.0198 | -35.0976 | -27.5651
18.0259 | -35.098 | -27.5724
18.032 | -35.0985 | -27.5792
18.0381 | -35.099 | -27.5861
18.0442 | -35.0995 | -27.5933
18.0503 | -35.0999 | -27.6006
18.0564 | -35.1004 | -27.6077
18.0625 | -35.1008 | -27.6147
18.0686 | -35.1013 | -27.6217
18.0747 | -35.1018 | -27.6287
18.0808 | -35.1023 | -27.6356
18.0869 | -35.1027 | -27.6427
18.093 | -35.1032 | -27.6497
18.0991 | -35.1037 | -27.6572
18.1052 | -35.1042 | -27.6648
18.1113 | -35.1046 | -27.6717
18.1174 | -35.1051 | -27.6786
18.1236 | -35.1056 | -27.6858
18.1297 | -35.106 | -27.693
18.1358 | -35.1066 | -27.7002
18.1419 | -35.1071 | -27.7075
18.148 | -35.1075 | -27.7149
18.1541 | -35.108 | -27.7217
18.1602 | -35.1085 | -27.729
18.1663 | -35.109 | -27.7363
18.1724 | -35.1095 | -27.7437
18.1785 | -35.1099 | -27.7509
18.1846 | -35.1104 | -27.7576
18.1907 | -35.1108 | -27.7651
18.1968 | -35.1113 | -27.7724
18.2029 | -35.1118 | -27.7792
18.209 | -35.1123 | -27.7864
18.2151 | -35.1128 | -27.7936
18.2212 | -35.1133 | -27.8004
18.2273 | -35.1138 | -27.8076
18.2334 | -35.1142 | -27.8146
18.2395 | -35.1147 | -27.8216
18.2456 | -35.1152 | -27.8287
18.2517 | -35.1156 | -27.8358
18.2578 | -35.1161 | -27.8427
18.2639 | -35.1166 | -27.85
18.2701 | -35.1171 | -27.8572
18.2762 | -35.1175 | -27.8646
18.2823 | -35.118 | -27.8719
18.2884 | -35.1185 | -27.8792
18.2945 | -35.119 | -27.8862
18.3006 | -35.1195 | -27.8933
18.3067 | -35.12 | -27.9002
18.3128 | -35.1205 | -27.9071
18.3189 | -35.121 | -27.9145
18.325 | -35.1214 | -27.9213
18.3311 | -35.1219 | -27.9287
18.3372 | -35.1224 | -27.9355
18.3433 | -35.1228 | -27.9423
18.3494 | -35.1232 | -27.9494
18.3555 | -35.1237 | -27.9562
18.3616 | -35.1242 | -27.9632
18.3677 | -35.1247 | -27.9698
18.3738 | -35.1251 | -27.9768
18.3799 | -35.1256 | -27.9837
18.386 | -35.126 | -27.9907
18.3921 | -35.1266 | -27.9979
18.3982 | -35.127 | -28.0049
18.4043 | -35.1275 | -28.012
18.4105 | -35.1279 | -28.019
18.4166 | -35.1284 | -28.0261
18.4227 | -35.1289 | -28.0335
18.4288 | -35.1294 | -28.0409
18.4349 | -35.1299 | -28.0479
18.441 | -35.1304 | -28.0549
18.4471 | -35.1309 | -28.0622
18.4532 | -35.1314 | -28.0694
18.4593 | -35.1319 | -28.0763
18.4654 | -35.1324 | -28.0832
18.4715 | -35.1329 | -28.0903
18.4776 | -35.1334 | -28.0974
18.4837 | -35.1339 | -28.1044
18.4898 | -35.1344 | -28.1116
18.4959 | -35.1349 | -28.1186
18.502 | -35.1355 | -28.1256
18.5081 | -35.136 | -28.1328
18.5142 | -35.1364 | -28.1403
18.5203 | -35.1369 | -28.147
18.5264 | -35.1373 | -28.1538
18.5325 | -35.1379 | -28.1609
18.5386 | -35.1384 | -28.1679
18.5447 | -35.1389 | -28.1748
18.5508 | -35.1394 | -28.1819
18.557 | -35.1399 | -28.1891
18.5631 | -35.1403 | -28.1962
18.5692 | -35.1408 | -28.2033
18.5753 | -35.1413 | -28.2101
18.5814 | -35.1418 | -28.2176
18.5875 | -35.1423 | -28.225
18.5936 | -35.1428 | -28.2324
18.5997 | -35.1433 | -28.2393
18.6058 | -35.1438 | -28.2463
18.6119 | -35.1443 | -28.2534
18.618 | -35.1448 | -28.2606
18.6241 | -35.1453 | -28.2679
18.6302 | -35.1458 | -28.2751
18.6363 | -35.1463 | -28.2818
18.6424 | -35.1468 | -28.2888
18.6485 | -35.1473 | -28.296
18.6546 | -35.1478 | -28.3031
18.6607 | -35.1482 | -28.31
18.6668 | -35.1487 | -28.3171
18.6729 | -35.1492 | -28.3243
18.679 | -35.1497 | -28.331
18.6851 | -35.1502 | -28.3382
18.6912 | -35.1507 | -28.3453
18.6974 | -35.1512 | -28.3523
18.7035 | -35.1516 | -28.3591
18.7096 | -35.1522 | -28.3662
18.7157 | -35.1527 | -28.3726
18.7218 | -35.1532 | -28.3795
18.7279 | -35.1537 | -28.3862
18.734 | -35.1541 | -28.3934
18.7401 | -35.1546 | -28.4002
18.7462 | -35.1551 | -28.407
18.7523 | -35.1556 | -28.414
18.7584 | -35.1561 | -28.4213
18.7645 | -35.1567 | -28.4278
18.7706 | -35.1572 | -28.4351
18.7767 | -35.1577 | -28.4424
18.7828 | -35.1582 | -28.4496
18.7889 | -35.1587 | -28.4567
18.795 | -35.1591 | -28.4637
18.8011 | -35.1596 | -28.4708
18.8072 | -35.1601 | -28.4779
18.8133 | -35.1606 | -28.4849
18.8194 | -35.1611 | -28.4914
18.8255 | -35.1616 | -28.4981
18.8316 | -35.1621 | -28.5052
18.8377 | -35.1625 | -28.5123
18.8439 | -35.163 | -28.5192
18.85 | -35.1635 | -28.5258
18.8561 | -35.164 | -28.5326
18.8622 | -35.1644 | -28.5399
18.8683 | -35.1649 | -28.5467
18.8744 | -35.1655 | -28.5534
18.8805 | -35.166 | -28.5604
18.8866 | -35.1665 | -28.5672
18.8927 | -35.167 | -28.5739
18.8988 | -35.1675 | -28.5808
18.9049 | -35.168 | -28.5879
18.911 | -35.1685 | -28.5948
18.9171 | -35.169 | -28.6019
18.9232 | -35.1695 | -28.6087
18.9293 | -35.17 | -28.6158
18.9354 | -35.1704 | -28.6225
18.9415 | -35.1709 | -28.6292
18.9476 | -35.1714 | -28.6364
18.9537 | -35.1719 | -28.6436
18.9598 | -35.1725 | -28.6501
18.9659 | -35.173 | -28.6573
18.972 | -35.1734 | -28.6642
18.9781 | -35.1739 | -28.671
18.9843 | -35.1745 | -28.6782
18.9904 | -35.1749 | -28.685
18.9965 | -35.1754 | -28.6918
19.0026 | -35.1759 | -28.6987
19.0087 | -35.1764 | -28.7056
19.0148 | -35.1769 | -28.7123
19.0209 | -35.1774 | -28.7193
19.027 | -35.1779 | -28.7263
19.0331 | -35.1784 | -28.7334
19.0392 | -35.1789 | -28.7401
19.0453 | -35.1794 | -28.7469
19.0514 | -35.1799 | -28.7539
19.0575 | -35.1804 | -28.7606
19.0636 | -35.1809 | -28.7671
19.0697 | -35.1814 | -28.7744
19.0758 | -35.182 | -28.7814
19.0819 | -35.1825 | -28.7885
19.088 | -35.183 | -28.7951
19.0941 | -35.1835 | -28.8018
19.1002 | -35.184 | -28.8083
19.1063 | -35.1846 | -28.8154
19.1124 | -35.1851 | -28.8221
19.1185 | -35.1856 | -28.8292
19.1246 | -35.1861 | -28.8361
19.1308 | -35.1866 | -28.8422
19.1369 | -35.1872 | -28.8487
19.143 | -35.1876 | -28.8555
19.1491 | -35.1882 | -28.8623
19.1552 | -35.1887 | -28.8693
19.1613 | -35.1892 | -28.8762
19.1674 | -35.1896 | -28.883
19.1735 | -35.1902 | -28.8898
19.1796 | -35.1907 | -28.8967
19.1857 | -35.1912 | -28.9037
19.1918 | -35.1917 | -28.9106
19.1979 | -35.1921 | -28.9175
19.204 | -35.1927 | -28.9242
19.2101 | -35.1932 | -28.9313
19.2162 | -35.1936 | -28.9384
19.2223 | -35.1941 | -28.9447
19.2284 | -35.1947 | -28.9519
19.2345 | -35.1952 | -28.9589
19.2406 | -35.1958 | -28.9655
19.2467 | -35.1962 | -28.9723
19.2528 | -35.1968 | -28.9791
19.2589 | -35.1972 | -28.9859
19.265 | -35.1978 | -28.9928
19.2712 | -35.1983 | -28.9997
19.2773 | -35.1988 | -29.0065
19.2834 | -35.1992 | -29.0134
19.2895 | -35.1997 | -29.0203
19.2956 | -35.2003 | -29.0272
19.3017 | -35.2008 | -29.0339
19.3078 | -35.2013 | -29.041
19.3139 | -35.2019 | -29.0474
19.32 | -35.2024 | -29.054
19.3261 | -35.2029 | -29.0609
19.3322 | -35.2034 | -29.0675
19.3383 | -35.2039 | -29.074
19.3444 | -35.2045 | -29.0809
19.3505 | -35.205 | -29.0878
19.3566 | -35.2054 | -29.0946
19.3627 | -35.206 | -29.1017
19.3688 | -35.2065 | -29.1087
19.3749 | -35.207 | -29.1157
19.381 | -35.2075 | -29.1224
19.3871 | -35.208 | -29.1293
19.3932 | -35.2085 | -29.1364
19.3993 | -35.209 | -29.143
19.4054 | -35.2096 | -29.1498
19.4115 | -35.2101 | -29.1567
19.4177 | -35.2106 | -29.1634
19.4238 | -35.2111 | -29.17
19.4299 | -35.2117 | -29.1766
19.436 | -35.2122 | -29.1835
19.4421 | -35.2127 | -29.1901
19.4482 | -35.2132 | -29.197
19.4543 | -35.2138 | -29.2035
19.4604 | -35.2142 | -29.2104
19.4665 | -35.2148 | -29.2172
19.4726 | -35.2153 | -29.224
19.4787 | -35.2158 | -29.2309
19.4848 | -35.2163 | -29.2377
19.4909 | -35.2169 | -29.2445
19.497 | -35.2174 | -29.2514
19.5031 | -35.2179 | -29.2582
19.5092 | -35.2184 | -29.2651
19.5153 | -35.2189 | -29.2721
19.5214 | -35.2195 | -29.2785
19.5275 | -35.2199 | -29.2854
19.5336 | -35.2204 | -29.2922
19.5397 | -35.221 | -29.299
19.5458 | -35.2215 | -29.3058
19.5519 | -35.222 | -29.3126
19.5581 | -35.2225 | -29.3193
19.5642 | -35.223 | -29.3263
19.5703 | -35.2236 | -29.3332
19.5764 | -35.2241 | -29.34
19.5825 | -35.2246 | -29.347
19.5886 | -35.2251 | -29.3539
19.5947 | -35.2256 | -29.3605
19.6008 | -35.2261 | -29.3675
19.6069 | -35.2266 | -29.3742
19.613 | -35.2271 | -29.3807
19.6191 | -35.2276 | -29.3878
19.6252 | -35.2281 | -29.3947
19.6313 | -35.2286 | -29.4017
19.6374 | -35.2291 | -29.4084
19.6435 | -35.2296 | -29.4151
19.6496 | -35.2301 | -29.4217
19.6557 | -35.2307 | -29.4284
19.6618 | -35.2312 | -29.4351
19.6679 | -35.2317 | -29.4418
19.674 | -35.2322 | -29.4485
19.6801 | -35.2327 | -29.4552
19.6862 | -35.2332 | -29.4621
19.6923 | -35.2337 | -29.4687
19.6984 | -35.2342 | -29.4751
19.7046 | -35.2348 | -29.482
19.7107 | -35.2353 | -29.4887
19.7168 | -35.2357 | -29.4956
19.7229 | -35.2362 | -29.5024
19.729 | -35.2367 | -29.5093
19.7351 | -35.2372 | -29.5155
19.7412 | -35.2378 | -29.5221
19.7473 | -35.2383 | -29.5288
19.7534 | -35.2388 | -29.5357
19.7595 | -35.2393 | -29.5427
19.7656 | -35.2399 | -29.5498
19.7717 | -35.2404 | -29.5563
19.7778 | -35.2409 | -29.563
19.7839 | -35.2414 | -29.5697
19.79 | -35.2419 | -29.5761
19.7961 | -35.2424 | -29.5829
19.8022 | -35.2429 | -29.5898
19.8083 | -35.2434 | -29.5964
19.8144 | -35.2439 | -29.6033
19.8205 | -35.2444 | -29.6101
19.8266 | -35.2449 | -29.617
19.8327 | -35.2454 | -29.6239
19.8388 | -35.2459 | -29.631
19.845 | -35.2465 | -29.6375
19.8511 | -35.247 | -29.6444
19.8572 | -35.2475 | -29.6513
19.8633 | -35.248 | -29.658
19.8694 | -35.2486 | -29.665
19.8755 | -35.2492 | -29.6717
19.8816 | -35.2497 | -29.6785
19.8877 | -35.2502 | -29.685
19.8938 | -35.2507 | -29.6916
19.8999 | -35.2512 | -29.6984
19.906 | -35.2518 | -29.7052
19.9121 | -35.2523 | -29.7117
19.9182 | -35.2528 | -29.7183
19.9243 | -35.2534 | -29.7254
19.9304 | -35.2539 | -29.7327
19.9365 | -35.2544 | -29.7396
19.9426 | -35.2549 | -29.7459
19.9487 | -35.2555 | -29.7523
19.9548 | -35.256 | -29.7595
19.9609 | -35.2565 | -29.766
19.967 | -35.2571 | -29.7727
19.9731 | -35.2575 | -29.7793
19.9792 | -35.258 | -29.786
19.9853 | -35.2586 | -29.7929
19.9915 | -35.2591 | -29.7998
19.9976 | -35.2596 | -29.8064
20.0037 | -35.2602 | -29.8132
20.0098 | -35.2607 | -29.8198
20.0159 | -35.2613 | -29.8264
20.022 | -35.2618 | -29.8334
20.0281 | -35.2624 | -29.8402
20.0342 | -35.2628 | -29.8469
20.0403 | -35.2633 | -29.8535
20.0464 | -35.2639 | -29.8605
20.0525 | -35.2644 | -29.8673
20.0586 | -35.2649 | -29.8738
20.0647 | -35.2654 | -29.8806
20.0708 | -35.266 | -29.8875
20.0769 | -35.2665 | -29.8941
20.083 | -35.267 | -29.901
20.0891 | -35.2674 | -29.9074
20.0952 | -35.268 | -29.914
20.1013 | -35.2685 | -29.9207
20.1074 | -35.269 | -29.9275
20.1135 | -35.2696 | -29.9341
20.1196 | -35.2701 | -29.9407
20.1257 | -35.2706 | -29.9475
20.1319 | -35.2712 | -29.9541
20.138 | -35.2717 | -29.9606
20.1441 | -35.2722 | -29.9672
20.1502 | -35.2728 | -29.9743
20.1563 | -35.2732 | -29.9805
20.1624 | -35.2738 | -29.9874
20.1685 | -35.2743 | -29.9941
20.1746 | -35.2748 | -30.0012
20.1807 | -35.2754 | -30.0079
20.1868 | -35.2759 | -30.0145
20.1929 | -35.2765 | -30.0212
20.199 | -35.2769 | -30.028
20.2051 | -35.2775 | -30.0348
20.2112 | -35.278 | -30.0414
20.2173 | -35.2784 | -30.0483
20.2234 | -35.2789 | -30.0552
20.2295 | -35.2795 | -30.0619
20.2356 | -35.2801 | -30.0688
20.2417 | -35.2806 | -30.0755
20.2478 | -35.2811 | -30.0822
20.2539 | -35.2816 | -30.0892
20.26 | -35.2821 | -30.0959
20.2661 | -35.2826 | -30.1025
20.2723 | -35.283 | -30.1094
20.2784 | -35.2836 | -30.1162
20.2845 | -35.2841 | -30.1233
20.2906 | -35.2846 | -30.13
20.2967 | -35.2852 | -30.1365
20.3028 | -35.2857 | -30.1433
20.3089 | -35.2862 | -30.1499
20.315 | -35.2867 | -30.1569
20.3211 | -35.2872 | -30.1639
20.3272 | -35.2877 | -30.1703
20.3333 | -35.2882 | -30.1767
20.3394 | -35.2887 | -30.1835
20.3455 | -35.2892 | -30.1903
20.3516 | -35.2898 | -30.1969
20.3577 | -35.2903 | -30.2035
20.3638 | -35.2908 | -30.2103
20.3699 | -35.2913 | -30.217
20.376 | -35.2919 | -30.224
20.3821 | -35.2924 | -30.2306
20.3882 | -35.2929 | -30.2371
20.3943 | -35.2934 | -30.2441
20.4004 | -35.2939 | -30.2505
20.4065 | -35.2945 | -30.2572
20.4126 | -35.295 | -30.2637
20.4188 | -35.2955 | -30.2704
20.4249 | -35.296 | -30.2774
20.431 | -35.2965 | -30.2844
20.4371 | -35.2971 | -30.2907
20.4432 | -35.2976 | -30.2974
20.4493 | -35.2981 | -30.304
20.4554 | -35.2987 | -30.3106
20.4615 | -35.2992 | -30.3175
20.4676 | -35.2997 | -30.3241
20.4737 | -35.3003 | -30.3303
20.4798 | -35.3008 | -30.3369
20.4859 | -35.3014 | -30.3435
20.492 | -35.3019 | -30.3503
20.4981 | -35.3024 | -30.3571
20.5042 | -35.3029 | -30.3634
20.5103 | -35.3034 | -30.3699
20.5164 | -35.3039 | -30.3768
20.5225 | -35.3044 | -30.3838
20.5286 | -35.3049 | -30.391
20.5347 | -35.3054 | -30.3973
20.5408 | -35.306 | -30.404
20.5469 | -35.3065 | -30.4106
20.553 | -35.307 | -30.4172
20.5592 | -35.3075 | -30.4238
20.5653 | -35.308 | -30.4305
20.5714 | -35.3086 | -30.437
20.5775 | -35.3091 | -30.4433
20.5836 | -35.3096 | -30.4501
20.5897 | -35.3102 | -30.4567
20.5958 | -35.3107 | -30.4632
20.6019 | -35.3112 | -30.4696
20.608 | -35.3117 | -30.4763
20.6141 | -35.3122 | -30.4825
20.6202 | -35.3127 | -30.4894
20.6263 | -35.3132 | -30.4959
20.6324 | -35.3137 | -30.5024
20.6385 | -35.3142 | -30.5088
20.6446 | -35.3147 | -30.5154
20.6507 | -35.3153 | -30.5219
20.6568 | -35.3158 | -30.5284
20.6629 | -35.3163 | -30.5349
20.669 | -35.3168 | -30.5415
20.6751 | -35.3173 | -30.5481
20.6812 | -35.3179 | -30.5549
20.6873 | -35.3184 | -30.5617
20.6934 | -35.319 | -30.5684
20.6995 | -35.3194 | -30.5748
20.7057 | -35.32 | -30.5814
20.7118 | -35.3205 | -30.5876
20.7179 | -35.3209 | -30.5942
20.724 | -35.3215 | -30.6007
20.7301 | -35.322 | -30.6077
20.7362 | -35.3225 | -30.6143
20.7423 | -35.323 | -30.6209
20.7484 | -35.3235 | -30.6276
20.7545 | -35.3241 | -30.634
20.7606 | -35.3246 | -30.6406
20.7667 | -35.3251 | -30.6472
20.7728 | -35.3256 | -30.6536
20.7789 | -35.3262 | -30.6602
20.785 | -35.3266 | -30.6668
20.7911 | -35.3271 | -30.6732
20.7972 | -35.3276 | -30.6797
20.8033 | -35.3281 | -30.6858
20.8094 | -35.3286 | -30.6925
20.8155 | -35.3291 | -30.6992
20.8216 | -35.3296 | -30.706
20.8277 | -35.3301 | -30.7125
20.8338 | -35.3306 | -30.7192
20.8399 | -35.3311 | -30.7258
20.8461 | -35.3316 | -30.7326
20.8522 | -35.3322 | -30.739
20.8583 | -35.3327 | -30.7453
20.8644 | -35.3332 | -30.7517
20.8705 | -35.3337 | -30.7582
20.8766 | -35.3343 | -30.7649
20.8827 | -35.3348 | -30.7714
20.8888 | -35.3353 | -30.7779
20.8949 | -35.3358 | -30.7846
20.901 | -35.3363 | -30.7911
20.9071 | -35.3368 | -30.7975
20.9132 | -35.3373 | -30.8041
20.9193 | -35.3378 | -30.8106
20.9254 | -35.3383 | -30.8171
20.9315 | -35.3388 | -30.8233
20.9376 | -35.3393 | -30.83
20.9437 | -35.3398 | -30.8367
20.9498 | -35.3404 | -30.8431
20.9559 | -35.3409 | -30.8495
20.962 | -35.3414 | -30.8565
20.9681 | -35.3419 | -30.863
20.9742 | -35.3424 | -30.8696
20.9803 | -35.3429 | -30.8761
20.9864 | -35.3434 | -30.8827
20.9926 | -35.3439 | -30.8892
20.9987 | -35.3444 | -30.8955
21.0048 | -35.3449 | -30.9022
21.0109 | -35.3454 | -30.9084
21.017 | -35.3459 | -30.915
21.0231 | -35.3464 | -30.9213
21.0292 | -35.3469 | -30.928
21.0353 | -35.3474 | -30.9342
21.0414 | -35.3479 | -30.9408
21.0475 | -35.3484 | -30.9474
21.0536 | -35.3489 | -30.9538
21.0597 | -35.3495 | -30.9603
21.0658 | -35.3499 | -30.9666
21.0719 | -35.3504 | -30.9736
21.078 | -35.3509 | -30.9805
21.0841 | -35.3514 | -30.9868
21.0902 | -35.3519 | -30.9934
21.0963 | -35.3524 | -30.9996
21.1024 | -35.3529 | -31.006
21.1085 | -35.3535 | -31.0123
21.1146 | -35.354 | -31.0192
21.1207 | -35.3545 | -31.0255
21.1268 | -35.355 | -31.0317
21.133 | -35.3556 | -31.038
21.1391 | -35.3561 | -31.0443
21.1452 | -35.3566 | -31.0512
21.1513 | -35.3571 | -31.0575
21.1574 | -35.3576 | -31.0642
21.1635 | -35.358 | -31.0704
21.1696 | -35.3585 | -31.077
21.1757 | -35.359 | -31.0831
21.1818 | -35.3595 | -31.0898
21.1879 | -35.3601 | -31.0964
21.194 | -35.3606 | -31.1027
21.2001 | -35.3611 | -31.1088
21.2062 | -35.3615 | -31.1156
21.2123 | -35.362 | -31.1222
21.2184 | -35.3625 | -31.129
21.2245 | -35.363 | -31.1359
21.2306 | -35.3635 | -31.1424
21.2367 | -35.364 | -31.1487
21.2428 | -35.3645 | -31.1549
21.2489 | -35.365 | -31.1612
21.255 | -35.3655 | -31.1676
21.2611 | -35.366 | -31.1742
21.2672 | -35.3665 | -31.1807
21.2733 | -35.367 | -31.1869
21.2795 | -35.3676 | -31.1937
21.2856 | -35.3681 | -31.2004
21.2917 | -35.3685 | -31.207
21.2978 | -35.369 | -31.2135
21.3039 | -35.3695 | -31.2199
21.31 | -35.37 | -31.2261
21.3161 | -35.3705 | -31.2324
21.3222 | -35.3709 | -31.2388
21.3283 | -35.3714 | -31.2458
21.3344 | -35.3719 | -31.2525
21.3405 | -35.3723 | -31.2592
21.3466 | -35.3728 | -31.266
21.3527 | -35.3733 | -31.2725
21.3588 | -35.3738 | -31.2789
21.3649 | -35.3742 | -31.285
21.371 | -35.3748 | -31.2913
21.3771 | -35.3753 | -31.2977
21.3832 | -35.3757 | -31.304
21.3893 | -35.3762 | -31.3102
21.3954 | -35.3766 | -31.3165
21.4015 | -35.3771 | -31.3227
21.4076 | -35.3776 | -31.3291
21.4137 | -35.3781 | -31.3356
21.4199 | -35.3786 | -31.3422
21.426 | -35.379 | -31.3488
21.4321 | -35.3795 | -31.3555
21.4382 | -35.38 | -31.3617
21.4443 | -35.3805 | -31.368
21.4504 | -35.381 | -31.3746
21.4565 | -35.3815 | -31.3808
21.4626 | -35.382 | -31.3874
21.4687 | -35.3826 | -31.3939
21.4748 | -35.3831 | -31.4001
21.4809 | -35.3836 | -31.4065
21.487 | -35.3841 | -31.4129
21.4931 | -35.3846 | -31.4191
21.4992 | -35.385 | -31.4253
21.5053 | -35.3855 | -31.4322
21.5114 | -35.386 | -31.4382
21.5175 | -35.3865 | -31.4446
21.5236 | -35.3869 | -31.4513
21.5297 | -35.3875 | -31.4576
21.5358 | -35.3879 | -31.464
21.5419 | -35.3884 | -31.4708
21.548 | -35.3889 | -31.4776
21.5541 | -35.3894 | -31.484
21.5602 | -35.3899 | -31.4904
21.5664 | -35.3904 | -31.4965
21.5725 | -35.3908 | -31.5028
21.5786 | -35.3913 | -31.5092
21.5847 | -35.3917 | -31.5156
21.5908 | -35.3922 | -31.5222
21.5969 | -35.3927 | -31.5289
21.603 | -35.3932 | -31.5352
21.6091 | -35.3938 | -31.5419
21.6152 | -35.3943 | -31.5484
21.6213 | -35.3948 | -31.555
21.6274 | -35.3953 | -31.5614
21.6335 | -35.3958 | -31.5676
21.6396 | -35.3962 | -31.574
21.6457 | -35.3968 | -31.5808
21.6518 | -35.3972 | -31.5876
21.6579 | -35.3977 | -31.5938
21.664 | -35.3982 | -31.6002
21.6701 | -35.3987 | -31.6067
21.6762 | -35.3992 | -31.6134
21.6823 | -35.3997 | -31.6196
21.6884 | -35.4001 | -31.6261
21.6945 | -35.4006 | -31.6325
21.7006 | -35.4011 | -31.6391
21.7068 | -35.4016 | -31.6456
21.7129 | -35.4021 | -31.652
21.719 | -35.4026 | -31.6583
21.7251 | -35.4031 | -31.6648
21.7312 | -35.4036 | -31.6709
21.7373 | -35.4041 | -31.6774
21.7434 | -35.4045 | -31.6838
21.7495 | -35.405 | -31.6901
21.7556 | -35.4055 | -31.6967
21.7617 | -35.406 | -31.7031
21.7678 | -35.4065 | -31.7097
21.7739 | -35.4069 | -31.7161
21.78 | -35.4073 | -31.7223
21.7861 | -35.4078 | -31.7284
21.7922 | -35.4083 | -31.7348
21.7983 | -35.4088 | -31.7414
21.8044 | -35.4093 | -31.7481
21.8105 | -35.4097 | -31.7543
21.8166 | -35.4102 | -31.7606
21.8227 | -35.4106 | -31.767
21.8288 | -35.4112 | -31.7731
21.8349 | -35.4117 | -31.7793
21.841 | -35.4122 | -31.7861
21.8471 | -35.4126 | -31.793
21.8533 | -35.413 | -31.7993
21.8594 | -35.4136 | -31.8058
21.8655 | -35.4141 | -31.8124
21.8716 | -35.4145 | -31.8189
21.8777 | -35.415 | -31.8249
21.8838 | -35.4155 | -31.831
21.8899 | -35.4159 | -31.8375
21.896 | -35.4164 | -31.8439
21.9021 | -35.4168 | -31.8501
21.9082 | -35.4173 | -31.8568
21.9143 | -35.4177 | -31.863
21.9204 | -35.4181 | -31.8693
21.9265 | -35.4186 | -31.8754
21.9326 | -35.4191 | -31.882
21.9387 | -35.4196 | -31.8886
21.9448 | -35.4201 | -31.8949
21.9509 | -35.4205 | -31.9011
21.957 | -35.421 | -31.9074
21.9631 | -35.4215 | -31.9139
21.9692 | -35.422 | -31.9205
21.9753 | -35.4225 | -31.9268
21.9814 | -35.4229 | -31.9331
21.9875 | -35.4233 | -31.9391
21.9937 | -35.4238 | -31.9456
21.9998 | -35.4243 | -31.9521
22.0059 | -35.4248 | -31.9586
22.012 | -35.4252 | -31.9648
22.0181 | -35.4257 | -31.9712
22.0242 | -35.4261 | -31.9776
22.0303 | -35.4266 | -31.9841
22.0364 | -35.4271 | -31.9904
22.0425 | -35.4275 | -31.9967
22.0486 | -35.428 | -32.0033
22.0547 | -35.4284 | -32.0097
22.0608 | -35.4288 | -32.0163
22.0669 | -35.4293 | -32.0223
22.073 | -35.4297 | -32.0288
22.0791 | -35.4302 | -32.035
22.0852 | -35.4307 | -32.0413
22.0913 | -35.4312 | -32.0476
22.0974 | -35.4316 | -32.0538
22.1035 | -35.432 | -32.0601
22.1096 | -35.4325 | -32.0664
22.1157 | -35.433 | -32.0726
22.1218 | -35.4335 | -32.0787
22.1279 | -35.4339 | -32.0848
22.134 | -35.4344 | -32.0907
22.1402 | -35.4348 | -32.097
22.1463 | -35.4353 | -32.1034
22.1524 | -35.4357 | -32.1098
22.1585 | -35.4362 | -32.1162
22.1646 | -35.4366 | -32.1222
22.1707 | -35.4371 | -32.1287
22.1768 | -35.4376 | -32.1347
22.1829 | -35.438 | -32.141
22.189 | -35.4385 | -32.147
22.1951 | -35.4389 | -32.1531
22.2012 | -35.4394 | -32.1595
22.2073 | -35.4398 | -32.1659
22.2134 | -35.4403 | -32.1723
22.2195 | -35.4408 | -32.179
22.2256 | -35.4412 | -32.185
22.2317 | -35.4417 | -32.1913
22.2378 | -35.4422 | -32.1974
22.2439 | -35.4427 | -32.2037
22.25 | -35.4432 | -32.2101
22.2561 | -35.4437 | -32.2165
22.2622 | -35.444 | -32.2233
22.2683 | -35.4445 | -32.2291
22.2744 | -35.4449 | -32.2352
22.2806 | -35.4454 | -32.2413
22.2867 | -35.4458 | -32.2476
22.2928 | -35.4463 | -32.2538
22.2989 | -35.4467 | -32.26
22.305 | -35.4471 | -32.2657
22.3111 | -35.4476 | -32.2722
22.3172 | -35.448 | -32.2784
22.3233 | -35.4485 | -32.2848
22.3294 | -35.4489 | -32.2911
22.3355 | -35.4494 | -32.2971
22.3416 | -35.4499 | -32.3034
22.3477 | -35.4503 | -32.3094
22.3538 | -35.4508 | -32.3156
22.3599 | -35.4513 | -32.3219
22.366 | -35.4518 | -32.3283
22.3721 | -35.4522 | -32.3345
22.3782 | -35.4527 | -32.3411
22.3843 | -35.4532 | -32.3474
22.3904 | -35.4536 | -32.3537
22.3965 | -35.4541 | -32.3601
22.4026 | -35.4545 | -32.3659
22.4087 | -35.4548 | -32.3723
22.4148 | -35.4553 | -32.3788
22.4209 | -35.4557 | -32.3849
22.4271 | -35.4562 | -32.3909
22.4332 | -35.4566 | -32.3968
22.4393 | -35.4571 | -32.403
22.4454 | -35.4575 | -32.4089
22.4515 | -35.4579 | -32.4154
22.4576 | -35.4584 | -32.4216
22.4637 | -35.4589 | -32.4278
22.4698 | -35.4593 | -32.4339
22.4759 | -35.4597 | -32.4399
22.482 | -35.4602 | -32.4463
22.4881 | -35.4606 | -32.4527
22.4942 | -35.4611 | -32.459
22.5003 | -35.4615 | -32.4655
22.5064 | -35.4619 | -32.4715
22.5125 | -35.4623 | -32.4778
22.5186 | -35.4627 | -32.4841
22.5247 | -35.4632 | -32.4903
22.5308 | -35.4636 | -32.4967
22.5369 | -35.4641 | -32.5032
22.543 | -35.4645 | -32.5095
22.5491 | -35.465 | -32.5155
22.5552 | -35.4654 | -32.5217
22.5613 | -35.4659 | -32.5279
22.5675 | -35.4663 | -32.5341
22.5736 | -35.4668 | -32.5405
22.5797 | -35.4672 | -32.5466
22.5858 | -35.4676 | -32.5527
22.5919 | -35.4681 | -32.5588
22.598 | -35.4685 | -32.5653
22.6041 | -35.4689 | -32.5715
22.6102 | -35.4694 | -32.5775
22.6163 | -35.4698 | -32.5834
22.6224 | -35.4702 | -32.5898
22.6285 | -35.4707 | -32.5959
22.6346 | -35.4711 | -32.6019
22.6407 | -35.4715 | -32.6081
22.6468 | -35.472 | -32.6141
22.6529 | -35.4724 | -32.6204
22.659 | -35.4729 | -32.6265
22.6651 | -35.4733 | -32.6327
22.6712 | -35.4737 | -32.6388
22.6773 | -35.4742 | -32.6452
22.6834 | -35.4746 | -32.6513
22.6895 | -35.475 | -32.6574
22.6956 | -35.4755 | -32.6636
22.7017 | -35.4759 | -32.6699
22.7079 | -35.4763 | -32.6764
22.714 | -35.4768 | -32.6828
22.7201 | -35.4772 | -32.6888
22.7262 | -35.4777 | -32.6951
22.7323 | -35.478 | -32.7012
22.7384 | -35.4785 | -32.7075
22.7445 | -35.4789 | -32.7136
22.7506 | -35.4794 | -32.7198
22.7567 | -35.4798 | -32.7262
22.7628 | -35.4802 | -32.7325
22.7689 | -35.4807 | -32.7387
22.775 | -35.4811 | -32.7452
22.7811 | -35.4816 | -32.7512
22.7872 | -35.482 | -32.7574
22.7933 | -35.4824 | -32.7634
22.7994 | -35.4828 | -32.7695
22.8055 | -35.4832 | -32.7759
22.8116 | -35.4836 | -32.7823
22.8177 | -35.484 | -32.7885
22.8238 | -35.4844 | -32.7949
22.8299 | -35.4848 | -32.8009
22.836 | -35.4852 | -32.8073
22.8421 | -35.4857 | -32.8133
22.8482 | -35.4861 | -32.8195
22.8544 | -35.4865 | -32.8257
22.8605 | -35.487 | -32.8321
22.8666 | -35.4874 | -32.838
22.8727 | -35.4878 | -32.8443
22.8788 | -35.4882 | -32.8504
22.8849 | -35.4887 | -32.8566
22.891 | -35.4892 | -32.863
22.8971 | -35.4896 | -32.8693
22.9032 | -35.49 | -32.8753
22.9093 | -35.4904 | -32.8812
22.9154 | -35.4908 | -32.8876
22.9215 | -35.4913 | -32.8937
22.9276 | -35.4917 | -32.8997
22.9337 | -35.4921 | -32.9053
22.9398 | -35.4926 | -32.9116
22.9459 | -35.493 | -32.918
22.952 | -35.4934 | -32.9241
22.9581 | -35.4938 | -32.9302
22.9642 | -35.4943 | -32.9365
22.9703 | -35.4947 | -32.9424
22.9764 | -35.4951 | -32.9486
22.9825 | -35.4955 | -32.9546
22.9886 | -35.4959 | -32.9606
22.9948 | -35.4963 | -32.9669
23.0009 | -35.4967 | -32.9731
23.007 | -35.4971 | -32.9793
23.0131 | -35.4976 | -32.9852
23.0192 | -35.498 | -32.9912
23.0253 | -35.4984 | -32.9973
23.0314 | -35.4988 | -33.0036
23.0375 | -35.4993 | -33.0097
23.0436 | -35.4997 | -33.0157
23.0497 | -35.5002 | -33.0219
23.0558 | -35.5006 | -33.0279
23.0619 | -35.501 | -33.0339
23.068 | -35.5014 | -33.0403
23.0741 | -35.5018 | -33.046
23.0802 | -35.5022 | -33.0522
23.0863 | -35.5026 | -33.0586
23.0924 | -35.503 | -33.0647
23.0985 | -35.5033 | -33.0708
23.1046 | -35.5038 | -33.0769
23.1107 | -35.5042 | -33.0829
23.1168 | -35.5046 | -33.0892
23.1229 | -35.505 | -33.0953
23.129 | -35.5053 | -33.1014
23.1351 | -35.5057 | -33.1078
23.1413 | -35.5062 | -33.114
23.1474 | -35.5066 | -33.12
23.1535 | -35.5071 | -33.1262
23.1596 | -35.5074 | -33.1321
23.1657 | -35.5078 | -33.138
23.1718 | -35.5083 | -33.1443
23.1779 | -35.5087 | -33.1505
23.184 | -35.5091 | -33.1565
23.1901 | -35.5095 | -33.1627
23.1962 | -35.5099 | -33.1687
23.2023 | -35.5103 | -33.1745
23.2084 | -35.5108 | -33.1809
23.2145 | -35.5112 | -33.1871
23.2206 | -35.5116 | -33.1931
23.2267 | -35.512 | -33.1993
23.2328 | -35.5124 | -33.2053
23.2389 | -35.5128 | -33.2117
23.245 | -35.5132 | -33.2178
23.2511 | -35.5136 | -33.2235
23.2572 | -35.514 | -33.2295
23.2633 | -35.5144 | -33.2357
23.2694 | -35.5148 | -33.2418
23.2755 | -35.5152 | -33.248
23.2817 | -35.5157 | -33.2539
23.2878 | -35.5161 | -33.2599
23.2939 | -35.5165 | -33.2663
23.3 | -35.5169 | -33.2723
23.3061 | -35.5173 | -33.2782
23.3122 | -35.5177 | -33.2841
23.3183 | -35.518 | -33.2902
23.3244 | -35.5185 | -33.2957
23.3305 | -35.5189 | -33.302
23.3366 | -35.5193 | -33.3083
23.3427 | -35.5197 | -33.3139
23.3488 | -35.5201 | -33.3201
23.3549 | -35.5205 | -33.3263
23.361 | -35.5209 | -33.3321
23.3671 | -35.5214 | -33.3383
23.3732 | -35.5218 | -33.3443
23.3793 | -35.5222 | -33.3503
23.3854 | -35.5226 | -33.3566
23.3915 | -35.523 | -33.3625
23.3976 | -35.5234 | -33.3685
23.4037 | -35.5238 | -33.3745
23.4098 | -35.5242 | -33.3806
23.4159 | -35.5246 | -33.3864
23.422 | -35.525 | -33.3924
23.4282 | -35.5254 | -33.3984
23.4343 | -35.5259 | -33.4043
23.4404 | -35.5263 | -33.4108
23.4465 | -35.5266 | -33.4169
23.4526 | -35.527 | -33.4232
23.4587 | -35.5274 | -33.4293
23.4648 | -35.5278 | -33.435
23.4709 | -35.5281 | -33.4411
23.477 | -35.5285 | -33.4475
23.4831 | -35.5289 | -33.4536
23.4892 | -35.5293 | -33.4598
23.4953 | -35.5297 | -33.4659
23.5014 | -35.5301 | -33.4716
23.5075 | -35.5305 | -33.4777
23.5136 | -35.5309 | -33.4837
23.5197 | -35.5313 | -33.4898
23.5258 | -35.5317 | -33.4959
23.5319 | -35.532 | -33.5017
23.538 | -35.5324 | -33.5079
23.5441 | -35.5328 | -33.5139
23.5502 | -35.5332 | -33.5203
23.5563 | -35.5336 | -33.5263
23.5624 | -35.534 | -33.5325
23.5686 | -35.5344 | -33.5384
23.5747 | -35.5348 | -33.5442
23.5808 | -35.5351 | -33.5502
23.5869 | -35.5355 | -33.5562
23.593 | -35.5359 | -33.5623
23.5991 | -35.5363 | -33.5682
23.6052 | -35.5368 | -33.5739
23.6113 | -35.5372 | -33.5798
23.6174 | -35.5375 | -33.5859
23.6235 | -35.5379 | -33.5919
23.6296 | -35.5383 | -33.5981
23.6357 | -35.5386 | -33.6045
23.6418 | -35.5391 | -33.6098
23.6479 | -35.5395 | -33.6161
23.654 | -35.5398 | -33.622
23.6601 | -35.5402 | -33.6279
23.6662 | -35.5406 | -33.6342
23.6723 | -35.541 | -33.6399
23.6784 | -35.5414 | -33.6458
23.6845 | -35.5418 | -33.6521
23.6906 | -35.5422 | -33.6582
23.6967 | -35.5425 | -33.664
23.7028 | -35.5429 | -33.6702
23.7089 | -35.5433 | -33.6766
23.7151 | -35.5437 | -33.6827
23.7212 | -35.5441 | -33.6886
23.7273 | -35.5446 | -33.6944
23.7334 | -35.545 | -33.7007
23.7395 | -35.5454 | -33.7066
23.7456 | -35.5457 | -33.7129
23.7517 | -35.5462 | -33.7192
23.7578 | -35.5465 | -33.7255
23.7639 | -35.5469 | -33.7317
23.77 | -35.5473 | -33.7376
23.7761 | -35.5477 | -33.7434
23.7822 | -35.5481 | -33.7494
23.7883 | -35.5484 | -33.7556
23.7944 | -35.5488 | -33.7619
23.8005 | -35.5492 | -33.768
23.8066 | -35.5496 | -33.7741
23.8127 | -35.5501 | -33.78
23.8188 | -35.5505 | -33.7859
23.8249 | -35.5509 | -33.792
23.831 | -35.5512 | -33.7981
23.8371 | -35.5516 | -33.8038
23.8432 | -35.5521 | -33.8099
23.8493 | -35.5525 | -33.8159
23.8555 | -35.5529 | -33.8221
23.8616 | -35.5533 | -33.8283
23.8677 | -35.5537 | -33.8343
23.8738 | -35.5541 | -33.8405
23.8799 | -35.5544 | -33.8461
23.886 | -35.5548 | -33.8518
23.8921 | -35.5552 | -33.8581
23.8982 | -35.5556 | -33.8643
23.9043 | -35.556 | -33.8705
23.9104 | -35.5563 | -33.8766
23.9165 | -35.5567 | -33.8823
23.9226 | -35.5571 | -33.888
23.9287 | -35.5575 | -33.894
23.9348 | -35.5579 | -33.8999
23.9409 | -35.5583 | -33.9056
23.947 | -35.5586 | -33.9118
23.9531 | -35.559 | -33.9174
23.9592 | -35.5594 | -33.9239
23.9653 | -35.5598 | -33.9298
23.9714 | -35.5602 | -33.936
23.9775 | -35.5605 | -33.9423
23.9836 | -35.561 | -33.9482
23.9897 | -35.5613 | -33.9539
23.9958 | -35.5617 | -33.9601
24.002 | -35.5621 | -33.9662
24.0081 | -35.5625 | -33.972
24.0142 | -35.5629 | -33.9781
24.0203 | -35.5633 | -33.984
24.0264 | -35.5636 | -33.9897
24.0325 | -35.564 | -33.9956
24.0386 | -35.5644 | -34.0016
24.0447 | -35.5648 | -34.0075
24.0508 | -35.5652 | -34.0133
24.0569 | -35.5656 | -34.0191
24.063 | -35.5659 | -34.0248
24.0691 | -35.5663 | -34.0308
24.0752 | -35.5667 | -34.0368
24.0813 | -35.5671 | -34.0428
24.0874 | -35.5675 | -34.0489
24.0935 | -35.5678 | -34.0549
24.0996 | -35.5682 | -34.0607
24.1057 | -35.5686 | -34.0666
24.1118 | -35.5689 | -34.0727
24.1179 | -35.5693 | -34.0787
24.124 | -35.5697 | -34.0844
24.1301 | -35.57 | -34.0902
24.1362 | -35.5705 | -34.0963
24.1424 | -35.5708 | -34.102
24.1485 | -35.5713 | -34.1075
24.1546 | -35.5716 | -34.1135
24.1607 | -35.572 | -34.1194
24.1668 | -35.5724 | -34.1256
24.1729 | -35.5728 | -34.1314
24.179 | -35.5732 | -34.1375
24.1851 | -35.5736 | -34.1436
24.1912 | -35.574 | -34.1498
24.1973 | -35.5743 | -34.1561
24.2034 | -35.5747 | -34.1614
24.2095 | -35.5752 | -34.1675
24.2156 | -35.5756 | -34.1735
24.2217 | -35.5759 | -34.1793
24.2278 | -35.5763 | -34.1852
24.2339 | -35.5767 | -34.1911
24.24 | -35.5771 | -34.1969
24.2461 | -35.5774 | -34.2029
24.2522 | -35.5778 | -34.2087
24.2583 | -35.5782 | -34.215
24.2644 | -35.5786 | -34.221
24.2705 | -35.579 | -34.2266
24.2766 | -35.5794 | -34.2328
24.2827 | -35.5798 | -34.2386
24.2889 | -35.5801 | -34.2444
24.295 | -35.5805 | -34.2504
24.3011 | -35.5809 | -34.2563
24.3072 | -35.5812 | -34.2619
24.3133 | -35.5816 | -34.2676
24.3194 | -35.582 | -34.2737
24.3255 | -35.5824 | -34.2796
24.3316 | -35.5828 | -34.2857
24.3377 | -35.5832 | -34.2918
24.3438 | -35.5835 | -34.2976
24.3499 | -35.5838 | -34.3033
24.356 | -35.5842 | -34.3094
24.3621 | -35.5845 | -34.3152
24.3682 | -35.5849 | -34.321
24.3743 | -35.5853 | -34.3271
24.3804 | -35.5857 | -34.3326
24.3865 | -35.586 | -34.3385
24.3926 | -35.5864 | -34.3443
24.3987 | -35.5868 | -34.3501
24.4048 | -35.5872 | -34.3562
24.4109 | -35.5875 | -34.3618
24.417 | -35.5879 | -34.3674
24.4231 | -35.5882 | -34.3735
24.4293 | -35.5886 | -34.3793
24.4354 | -35.589 | -34.3848
24.4415 | -35.5893 | -34.3905
24.4476 | -35.5897 | -34.396
24.4537 | -35.5902 | -34.4018
24.4598 | -35.5906 | -34.4078
24.4659 | -35.5909 | -34.4139
24.472 | -35.5913 | -34.4199
24.4781 | -35.5916 | -34.4258
24.4842 | -35.592 | -34.4313
24.4903 | -35.5923 | -34.4373
24.4964 | -35.5927 | -34.4433
24.5025 | -35.5931 | -34.4493
24.5086 | -35.5934 | -34.455
24.5147 | -35.5939 | -34.4608
24.5208 | -35.5942 | -34.4664
24.5269 | -35.5946 | -34.472
24.533 | -35.595 | -34.4778
24.5391 | -35.5954 | -34.4839
24.5452 | -35.5958 | -34.4896
24.5513 | -35.5961 | -34.4952
24.5574 | -35.5965 | -34.5009
24.5635 | -35.5969 | -34.5065
24.5696 | -35.5973 | -34.5126
24.5758 | -35.5977 | -34.5186
24.5819 | -35.5981 | -34.5247
24.588 | -35.5984 | -34.5303
24.5941 | -35.5988 | -34.536
24.6002 | -35.5991 | -34.542
24.6063 | -35.5995 | -34.5477
24.6124 | -35.5998 | -34.5538
24.6185 | -35.6001 | -34.5597
24.6246 | -35.6005 | -34.5656
24.6307 | -35.6009 | -34.5715
24.6368 | -35.6013 | -34.5772
24.6429 | -35.6016 | -34.583
24.649 | -35.602 | -34.5887
24.6551 | -35.6023 | -34.5945
24.6612 | -35.6027 | -34.5999
24.6673 | -35.6031 | -34.6056
24.6734 | -35.6034 | -34.6114
24.6795 | -35.6039 | -34.6172
24.6856 | -35.6043 | -34.623
24.6917 | -35.6046 | -34.6289
24.6978 | -35.605 | -34.6344
24.7039 | -35.6054 | -34.6402
24.71 | -35.6058 | -34.6461
24.7162 | -35.6061 | -34.6517
24.7223 | -35.6065 | -34.6574
24.7284 | -35.6068 | -34.663
24.7345 | -35.6072 | -34.669
24.7406 | -35.6076 | -34.6749
24.7467 | -35.608 | -34.6807
24.7528 | -35.6083 | -34.6864
24.7589 | -35.6087 | -34.6921
24.765 | -35.609 | -34.6976
24.7711 | -35.6094 | -34.7033
24.7772 | -35.6097 | -34.709
24.7833 | -35.61 | -34.7151
24.7894 | -35.6104 | -34.7211
24.7955 | -35.6107 | -34.7273
24.8016 | -35.6111 | -34.7331
24.8077 | -35.6115 | -34.7389
24.8138 | -35.6119 | -34.7447
24.8199 | -35.6123 | -34.7504
24.826 | -35.6126 | -34.7559
24.8321 | -35.613 | -34.7619
24.8382 | -35.6134 | -34.7676
24.8443 | -35.6138 | -34.7736
24.8504 | -35.6141 | -34.7791
24.8565 | -35.6145 | -34.7848
24.8627 | -35.6149 | -34.7905
24.8688 | -35.6152 | -34.7959
24.8749 | -35.6156 | -34.8016
24.881 | -35.616 | -34.8076
24.8871 | -35.6164 | -34.8132
24.8932 | -35.6168 | -34.8189
24.8993 | -35.6171 | -34.8247
24.9054 | -35.6175 | -34.8302
24.9115 | -35.6179 | -34.8361
24.9176 | -35.6183 | -34.8418
24.9237 | -35.6187 | -34.8477
24.9298 | -35.619 | -34.8536
24.9359 | -35.6194 | -34.8594
24.942 | -35.6198 | -34.8649
24.9481 | -35.6202 | -34.8706
24.9542 | -35.6206 | -34.8766
24.9603 | -35.6209 | -34.8823
24.9664 | -35.6213 | -34.8881
24.9725 | -35.6217 | -34.8938
24.9786 | -35.6221 | -34.8993
24.9847 | -35.6224 | -34.9053
24.9908 | -35.6228 | -34.9112
24.9969 | -35.6232 | -34.9172
25.0031 | -35.6235 | -34.9228
25.0092 | -35.6239 | -34.9284
25.0153 | -35.6243 | -34.9343
25.0214 | -35.6246 | -34.9401
25.0275 | -35.625 | -34.9459
25.0336 | -35.6253 | -34.9517
25.0397 | -35.6257 | -34.9577
25.0458 | -35.626 | -34.9634
25.0519 | -35.6264 | -34.9689
25.058 | -35.6267 | -34.9748
25.0641 | -35.6271 | -34.9809
25.0702 | -35.6275 | -34.9865
25.0763 | -35.6278 | -34.9921
25.0824 | -35.6282 | -34.9977
25.0885 | -35.6285 | -35.0034
25.0946 | -35.6289 | -35.0091
25.1007 | -35.6293 | -35.0148
25.1068 | -35.6297 | -35.0202
25.1129 | -35.63 | -35.0261
25.119 | -35.6303 | -35.0314
25.1251 | -35.6306 | -35.0373
25.1312 | -35.6309 | -35.043
25.1373 | -35.6313 | -35.0486
25.1435 | -35.6317 | -35.0545
25.1496 | -35.6321 | -35.0599
25.1557 | -35.6324 | -35.0655
25.1618 | -35.6328 | -35.0712
25.1679 | -35.6331 | -35.077
25.174 | -35.6335 | -35.0829
25.1801 | -35.6339 | -35.0887
25.1862 | -35.6343 | -35.0943
25.1923 | -35.6347 | -35.0999
25.1984 | -35.635 | -35.1056
25.2045 | -35.6354 | -35.1113
25.2106 | -35.6358 | -35.1171
25.2167 | -35.6361 | -35.1229
25.2228 | -35.6366 | -35.1287
25.2289 | -35.6369 | -35.1342
25.235 | -35.6373 | -35.1398
25.2411 | -35.6376 | -35.1455
25.2472 | -35.638 | -35.1511
25.2533 | -35.6383 | -35.1569
25.2594 | -35.6386 | -35.1622
25.2655 | -35.639 | -35.168
25.2716 | -35.6393 | -35.1738
25.2777 | -35.6397 | -35.1796
25.2838 | -35.64 | -35.1851
25.29 | -35.6404 | -35.1909
25.2961 | -35.6407 | -35.1963
25.3022 | -35.641 | -35.2022
25.3083 | -35.6414 | -35.2077
25.3144 | -35.6418 | -35.2135
25.3205 | -35.6421 | -35.219
25.3266 | -35.6425 | -35.2247
25.3327 | -35.6428 | -35.2303
25.3388 | -35.6432 | -35.2359
25.3449 | -35.6436 | -35.2419
25.351 | -35.6439 | -35.2476
25.3571 | -35.6442 | -35.2533
25.3632 | -35.6445 | -35.2586
25.3693 | -35.6449 | -35.2643
25.3754 | -35.6452 | -35.27
25.3815 | -35.6456 | -35.2756
25.3876 | -35.646 | -35.2814
25.3937 | -35.6463 | -35.2872
25.3998 | -35.6467 | -35.2929
25.4059 | -35.647 | -35.2987
25.412 | -35.6474 | -35.3043
25.4181 | -35.6477 | -35.3103
25.4242 | -35.6481 | -35.3159
25.4304 | -35.6485 | -35.3216
25.4365 | -35.6488 | -35.3271
25.4426 | -35.6492 | -35.3328
25.4487 | -35.6496 | -35.3386
25.4548 | -35.6499 | -35.3444
25.4609 | -35.6503 | -35.3499
25.467 | -35.6506 | -35.3554
25.4731 | -35.651 | -35.3612
25.4792 | -35.6513 | -35.367
25.4853 | -35.6516 | -35.3724
25.4914 | -35.652 | -35.378
25.4975 | -35.6524 | -35.3834
25.5036 | -35.6528 | -35.389
25.5097 | -35.6531 | -35.3946
25.5158 | -35.6535 | -35.4002
25.5219 | -35.6538 | -35.4056
25.528 | -35.6541 | -35.4111
25.5341 | -35.6545 | -35.4168
25.5402 | -35.6548 | -35.4224
25.5463 | -35.6552 | -35.4282
25.5524 | -35.6555 | -35.4342
25.5585 | -35.6559 | -35.4396
25.5646 | -35.6563 | -35.4452
25.5707 | -35.6566 | -35.4507
25.5769 | -35.6569 | -35.456
25.583 | -35.6573 | -35.4618
25.5891 | -35.6577 | -35.4676
25.5952 | -35.658 | -35.473
25.6013 | -35.6584 | -35.4785
25.6074 | -35.6588 | -35.4839
25.6135 | -35.6591 | -35.4894
25.6196 | -35.6595 | -35.4953
25.6257 | -35.6598 | -35.5009
25.6318 | -35.6602 | -35.5067
25.6379 | -35.6605 | -35.5122
25.644 | -35.6608 | -35.5175
25.6501 | -35.6612 | -35.523
25.6562 | -35.6615 | -35.5287
25.6623 | -35.6619 | -35.5346
25.6684 | -35.6623 | -35.5404
25.6745 | -35.6626 | -35.5463
25.6806 | -35.6629 | -35.5512
25.6867 | -35.6633 | -35.5565
25.6928 | -35.6637 | -35.5621
25.6989 | -35.6641 | -35.5675
25.705 | -35.6644 | -35.5732
25.7111 | -35.6647 | -35.5788
25.7173 | -35.665 | -35.5844
25.7234 | -35.6654 | -35.59
25.7295 | -35.6658 | -35.5956
25.7356 | -35.6662 | -35.6013
25.7417 | -35.6665 | -35.607
25.7478 | -35.6669 | -35.6122
25.7539 | -35.6673 | -35.6178
25.76 | -35.6676 | -35.6234
25.7661 | -35.6679 | -35.6288
25.7722 | -35.6682 | -35.6344
25.7783 | -35.6686 | -35.64
25.7844 | -35.669 | -35.6454
25.7905 | -35.6694 | -35.6513
25.7966 | -35.6697 | -35.6568
25.8027 | -35.67 | -35.6623
25.8088 | -35.6704 | -35.6677
25.8149 | -35.6708 | -35.6731
25.821 | -35.6711 | -35.6781
25.8271 | -35.6714 | -35.6841
25.8332 | -35.6718 | -35.6897
25.8393 | -35.6722 | -35.6954
25.8454 | -35.6726 | -35.7008
25.8515 | -35.6729 | -35.7065
25.8576 | -35.6733 | -35.7119
25.8638 | -35.6737 | -35.7174
25.8699 | -35.674 | -35.7231
25.876 | -35.6744 | -35.7287
25.8821 | -35.6747 | -35.7343
25.8882 | -35.6751 | -35.7397
25.8943 | -35.6754 | -35.7453
25.9004 | -35.6758 | -35.7506
25.9065 | -35.6761 | -35.7563
25.9126 | -35.6765 | -35.7619
25.9187 | -35.6768 | -35.7675
25.9248 | -35.6772 | -35.7732
25.9309 | -35.6775 | -35.7788
25.937 | -35.6779 | -35.7844
25.9431 | -35.6783 | -35.7901
25.9492 | -35.6786 | -35.7954
25.9553 | -35.6789 | -35.8008
25.9614 | -35.6793 | -35.8059
25.9675 | -35.6797 | -35.8116
25.9736 | -35.68 | -35.8171
25.9797 | -35.6804 | -35.8225
25.9858 | -35.6807 | -35.828
25.9919 | -35.681 | -35.8338
25.998 | -35.6814 | -35.839
26.0042 | -35.6817 | -35.8447
26.0103 | -35.682 | -35.8502
26.0164 | -35.6824 | -35.8557
26.0225 | -35.6827 | -35.8611
26.0286 | -35.6831 | -35.8668
26.0347 | -35.6834 | -35.8722
26.0408 | -35.6837 | -35.8779
26.0469 | -35.6841 | -35.8831
26.053 | -35.6844 | -35.8887
26.0591 | -35.6847 | -35.8945
26.0652 | -35.6851 | -35.8997
26.0713 | -35.6854 | -35.9055
26.0774 | -35.6858 | -35.9111
26.0835 | -35.6861 | -35.9165
26.0896 | -35.6864 | -35.9219
26.0957 | -35.6868 | -35.9274
26.1018 | -35.6871 | -35.9327
26.1079 | -35.6875 | -35.9381
26.114 | -35.6878 | -35.9441
26.1201 | -35.6881 | -35.9496
26.1262 | -35.6885 | -35.9551
26.1323 | -35.6888 | -35.9604
26.1384 | -35.6892 | -35.9655
26.1445 | -35.6896 | -35.9709
26.1507 | -35.6899 | -35.9764
26.1568 | -35.6903 | -35.9821
26.1629 | -35.6906 | -35.9875
26.169 | -35.6909 | -35.9929
26.1751 | -35.6913 | -35.9981
26.1812 | -35.6916 | -36.0036
26.1873 | -35.692 | -36.0092
26.1934 | -35.6923 | -36.0145
26.1995 | -35.6926 | -36.0201
26.2056 | -35.693 | -36.0255
26.2117 | -35.6933 | -36.0308
26.2178 | -35.6936 | -36.0362
26.2239 | -35.694 | -36.0417
26.23 | -35.6943 | -36.0473
26.2361 | -35.6947 | -36.0527
26.2422 | -35.695 | -36.058
26.2483 | -35.6954 | -36.0634
26.2544 | -35.6957 | -36.069
26.2605 | -35.6961 | -36.0747
26.2666 | -35.6964 | -36.0798
26.2727 | -35.6967 | -36.0852
26.2788 | -35.6971 | -36.0904
26.2849 | -35.6974 | -36.0958
26.2911 | -35.6976 | -36.1009
26.2972 | -35.698 | -36.1064
26.3033 | -35.6984 | -36.1115
26.3094 | -35.6987 | -36.1173
26.3155 | -35.699 | -36.1221
26.3216 | -35.6993 | -36.1276
26.3277 | -35.6997 | -36.133
26.3338 | -35.7 | -36.138
26.3399 | -35.7003 | -36.1433
26.346 | -35.7007 | -36.1483
26.3521 | -35.7011 | -36.1539
26.3582 | -35.7014 | -36.1593
26.3643 | -35.7018 | -36.1649
26.3704 | -35.7021 | -36.1706
26.3765 | -35.7025 | -36.1761
26.3826 | -35.7029 | -36.1813
26.3887 | -35.7032 | -36.1867
26.3948 | -35.7036 | -36.1921
26.4009 | -35.7039 | -36.1972
26.407 | -35.7043 | -36.2026
26.4131 | -35.7046 | -36.2079
26.4192 | -35.7049 | -36.2134
26.4253 | -35.7053 | -36.2186
26.4314 | -35.7056 | -36.2239
26.4376 | -35.706 | -36.2293
26.4437 | -35.7063 | -36.2347
26.4498 | -35.7067 | -36.2402
26.4559 | -35.707 | -36.2452
26.462 | -35.7073 | -36.2507
26.4681 | -35.7076 | -36.2561
26.4742 | -35.708 | -36.2618
26.4803 | -35.7084 | -36.2669
26.4864 | -35.7087 | -36.272
26.4925 | -35.709 | -36.2774
26.4986 | -35.7093 | -36.2828
26.5047 | -35.7097 | -36.2882
26.5108 | -35.71 | -36.2937
26.5169 | -35.7104 | -36.2991
26.523 | -35.7107 | -36.3042
26.5291 | -35.711 | -36.3092
26.5352 | -35.7113 | -36.3148
26.5413 | -35.7117 | -36.3203
26.5474 | -35.712 | -36.3254
26.5535 | -35.7124 | -36.3309
26.5596 | -35.7127 | -36.3362
26.5657 | -35.713 | -36.3416
26.5718 | -35.7133 | -36.347
26.578 | -35.7136 | -36.3525
26.5841 | -35.714 | -36.358
26.5902 | -35.7143 | -36.3636
26.5963 | -35.7146 | -36.3687
26.6024 | -35.7149 | -36.3742
26.6085 | -35.7153 | -36.3795
26.6146 | -35.7156 | -36.3848
26.6207 | -35.716 | -36.3904
26.6268 | -35.7164 | -36.3956
26.6329 | -35.7167 | -36.4009
26.639 | -35.717 | -36.4064
26.6451 | -35.7174 | -36.4116
26.6512 | -35.7176 | -36.4167
26.6573 | -35.7179 | -36.4224
26.6634 | -35.7182 | -36.4278
26.6695 | -35.7185 | -36.433
26.6756 | -35.7189 | -36.4382
26.6817 | -35.7192 | -36.4435
26.6878 | -35.7196 | -36.4487
26.6939 | -35.7199 | -36.454
26.7 | -35.7202 | -36.459
26.7061 | -35.7206 | -36.4645
26.7122 | -35.7209 | -36.47
26.7183 | -35.7213 | -36.4753
26.7245 | -35.7216 | -36.4808
26.7306 | -35.7219 | -36.4862
26.7367 | -35.7222 | -36.4912
26.7428 | -35.7226 | -36.4965
26.7489 | -35.7229 | -36.502
26.755 | -35.7232 | -36.5075
26.7611 | -35.7236 | -36.5128
26.7672 | -35.7239 | -36.5182
26.7733 | -35.7243 | -36.5233
26.7794 | -35.7246 | -36.5289
26.7855 | -35.7249 | -36.5347
26.7916 | -35.7253 | -36.5395
26.7977 | -35.7257 | -36.5448
26.8038 | -35.726 | -36.5499
26.8099 | -35.7263 | -36.5554
26.816 | -35.7266 | -36.5609
26.8221 | -35.7269 | -36.5662
26.8282 | -35.7273 | -36.5718
26.8343 | -35.7276 | -36.5772
26.8404 | -35.728 | -36.5823
26.8465 | -35.7283 | -36.5879
26.8526 | -35.7286 | -36.5933
26.8587 | -35.7289 | -36.5987
26.8649 | -35.7293 | -36.6039
26.871 | -35.7296 | -36.6092
26.8771 | -35.7299 | -36.6144
26.8832 | -35.7302 | -36.6198
26.8893 | -35.7305 | -36.6249
26.8954 | -35.7309 | -36.6303
26.9015 | -35.7313 | -36.6357
26.9076 | -35.7316 | -36.6413
26.9137 | -35.732 | -36.6462
26.9198 | -35.7324 | -36.6515
26.9259 | -35.7327 | -36.6567
26.932 | -35.733 | -36.6621
26.9381 | -35.7333 | -36.6677
26.9442 | -35.7336 | -36.6727
26.9503 | -35.7339 | -36.6781
26.9564 | -35.7343 | -36.6835
26.9625 | -35.7346 | -36.6887
26.9686 | -35.735 | -36.694
26.9747 | -35.7354 | -36.6995
26.9808 | -35.7357 | -36.7042
26.9869 | -35.736 | -36.7095
26.993 | -35.7363 | -36.7147
26.9991 | -35.7367 | -36.7197
27.0052 | -35.737 | -36.7254
27.0114 | -35.7373 | -36.7308
27.0175 | -35.7376 | -36.736
27.0236 | -35.738 | -36.7415
27.0297 | -35.7383 | -36.7468
27.0358 | -35.7386 | -36.7521
27.0419 | -35.7389 | -36.7572
27.048 | -35.7393 | -36.7626
27.0541 | -35.7396 | -36.7679
27.0602 | -35.7399 | -36.7734
27.0663 | -35.7402 | -36.7787
27.0724 | -35.7405 | -36.7839
27.0785 | -35.7409 | -36.7892
27.0846 | -35.7412 | -36.7941
27.0907 | -35.7415 | -36.7992
27.0968 | -35.7418 | -36.8045
27.1029 | -35.7422 | -36.8097
27.109 | -35.7425 | -36.815
27.1151 | -35.7428 | -36.8201
27.1212 | -35.7431 | -36.8249
27.1273 | -35.7434 | -36.8301
27.1334 | -35.7437 | -36.8353
27.1395 | -35.744 | -36.8407
27.1456 | -35.7443 | -36.8459
27.1518 | -35.7446 | -36.8513
27.1579 | -35.745 | -36.856
27.164 | -35.7453 | -36.8613
27.1701 | -35.7456 | -36.8668
27.1762 | -35.7459 | -36.8723
27.1823 | -35.7461 | -36.8777
27.1884 | -35.7465 | -36.8829
27.1945 | -35.7469 | -36.888
27.2006 | -35.7472 | -36.8932
27.2067 | -35.7475 | -36.8984
27.2128 | -35.7478 | -36.9038
27.2189 | -35.7481 | -36.9089
27.225 | -35.7485 | -36.9141
27.2311 | -35.7488 | -36.9193
27.2372 | -35.7491 | -36.9247
27.2433 | -35.7494 | -36.9298
27.2494 | -35.7497 | -36.9352
27.2555 | -35.75 | -36.9403
27.2616 | -35.7503 | -36.9454
27.2677 | -35.7507 | -36.9504
27.2738 | -35.751 | -36.9559
27.2799 | -35.7514 | -36.9608
27.286 | -35.7517 | -36.966
27.2921 | -35.752 | -36.9712
27.2983 | -35.7523 | -36.9764
27.3044 | -35.7527 | -36.9815
27.3105 | -35.753 | -36.9869
27.3166 | -35.7533 | -36.9921
27.3227 | -35.7536 | -36.997
27.3288 | -35.7539 | -37.0022
27.3349 | -35.7542 | -37.0076
27.341 | -35.7545 | -37.0131
27.3471 | -35.7548 | -37.0185
27.3532 | -35.7551 | -37.0237
27.3593 | -35.7555 | -37.0291
27.3654 | -35.7558 | -37.0344
27.3715 | -35.7561 | -37.0392
27.3776 | -35.7565 | -37.0446
27.3837 | -35.7568 | -37.0494
27.3898 | -35.7571 | -37.0547
27.3959 | -35.7575 | -37.0598
27.402 | -35.7578 | -37.0647
27.4081 | -35.7581 | -37.07
27.4142 | -35.7584 | -37.0753
27.4203 | -35.7587 | -37.0806
27.4264 | -35.759 | -37.0857
27.4325 | -35.7593 | -37.091
27.4387 | -35.7597 | -37.0959
27.4448 | -35.76 | -37.1013
27.4509 | -35.7603 | -37.1065
27.457 | -35.7606 | -37.1117
27.4631 | -35.7609 | -37.1169
27.4692 | -35.7613 | -37.1222
27.4753 | -35.7615 | -37.1269
27.4814 | -35.7618 | -37.1319
27.4875 | -35.7621 | -37.1372
27.4936 | -35.7624 | -37.1422
27.4997 | -35.7628 | -37.1475
27.5058 | -35.7631 | -37.1524
27.5119 | -35.7633 | -37.1573
27.518 | -35.7636 | -37.1626
27.5241 | -35.7639 | -37.1677
27.5302 | -35.7643 | -37.1729
27.5363 | -35.7646 | -37.178
27.5424 | -35.7649 | -37.1827
27.5485 | -35.7652 | -37.1876
27.5546 | -35.7655 | -37.1928
27.5607 | -35.7658 | -37.1979
27.5668 | -35.7661 | -37.203
27.5729 | -35.7664 | -37.2079
27.5791 | -35.7667 | -37.2126
27.5852 | -35.767 | -37.2178
27.5913 | -35.7673 | -37.223
27.5974 | -35.7676 | -37.2283
27.6035 | -35.7679 | -37.2336
27.6096 | -35.7682 | -37.2387
27.6157 | -35.7685 | -37.2436
27.6218 | -35.7688 | -37.2486
27.6279 | -35.7692 | -37.2538
27.634 | -35.7695 | -37.2589
27.6401 | -35.7698 | -37.2641
27.6462 | -35.7701 | -37.2689
27.6523 | -35.7705 | -37.2738
27.6584 | -35.7707 | -37.2792
27.6645 | -35.771 | -37.2841
27.6706 | -35.7713 | -37.2891
27.6767 | -35.7716 | -37.294
27.6828 | -35.7719 | -37.2988
27.6889 | -35.7722 | -37.3039
27.695 | -35.7725 | -37.3089
27.7011 | -35.7728 | -37.3141
27.7072 | -35.7731 | -37.3194
27.7133 | -35.7735 | -37.3242
27.7194 | -35.7738 | -37.3288
27.7256 | -35.7741 | -37.3337
27.7317 | -35.7744 | -37.339
27.7378 | -35.7747 | -37.3444
27.7439 | -35.7751 | -37.3493
27.75 | -35.7754 | -37.3545
27.7561 | -35.7757 | -37.3591
27.7622 | -35.776 | -37.3643
27.7683 | -35.7763 | -37.3695
27.7744 | -35.7766 | -37.3744
27.7805 | -35.7769 | -37.3797
27.7866 | -35.7772 | -37.3845
27.7927 | -35.7775 | -37.3899
27.7988 | -35.7778 | -37.3947
27.8049 | -35.7781 | -37.3997
27.811 | -35.7784 | -37.405
27.8171 | -35.7788 | -37.4099
27.8232 | -35.7791 | -37.4147
27.8293 | -35.7794 | -37.4198
27.8354 | -35.7797 | -37.425
27.8415 | -35.78 | -37.4298
27.8476 | -35.7803 | -37.4349
27.8537 | -35.7806 | -37.4398
27.8598 | -35.7809 | -37.4446
27.866 | -35.7813 | -37.4495
27.8721 | -35.7816 | -37.4545
27.8782 | -35.7819 | -37.4594
27.8843 | -35.7822 | -37.4644
27.8904 | -35.7825 | -37.4693
27.8965 | -35.7828 | -37.4742
27.9026 | -35.7831 | -37.4792
27.9087 | -35.7834 | -37.4843
27.9148 | -35.7838 | -37.4894
27.9209 | -35.7841 | -37.4944
27.927 | -35.7844 | -37.4994
27.9331 | -35.7847 | -37.5042
27.9392 | -35.785 | -37.5094
27.9453 | -35.7853 | -37.5145
27.9514 | -35.7856 | -37.5195
27.9575 | -35.786 | -37.5241
27.9636 | -35.7863 | -37.529
27.9697 | -35.7865 | -37.5337
27.9758 | -35.7868 | -37.5387
27.9819 | -35.7871 | -37.544
27.988 | -35.7874 | -37.5489
27.9941 | -35.7877 | -37.5539
28.0002 | -35.788 | -37.5587
28.0063 | -35.7883 | -37.5639
28.0125 | -35.7886 | -37.5688
28.0186 | -35.7888 | -37.5741
28.0247 | -35.7891 | -37.5788
28.0308 | -35.7894 | -37.5837
28.0369 | -35.7897 | -37.5885
28.043 | -35.79 | -37.5937
28.0491 | -35.7904 | -37.5988
28.0552 | -35.7906 | -37.604
28.0613 | -35.7909 | -37.609
28.0674 | -35.7912 | -37.6139
28.0735 | -35.7915 | -37.6183

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.2716 | -41.8486
21.4908 | -35.5524 | -34.6751
16.4948 | -34.9273 | -27.4046
12.6602 | -34.5684 | -21.1989
9.7171 | -34.3577 | -16.3386
7.4582 | -34.2121 | -12.3958
5.7244 | -34.1045 | -9.2009
4.3936 | -34.041 | -6.6219
3.3722 | -34.0176 | -4.5913
2.5883 | -34.0123 | -3.049
1.9866 | -34.0182 | -1.816
1.5248 | -34.0163 | -0.82888
1.1703 | -34.0326 | -0.18357
1 | -34.0265 | 0.13562
0.89822 | -34.0209 | 0.40589
0.68941 | -34.0088 | 0.98204
0.52919 | -34.0148 | 1.4285
0.40612 | -34.0277 | 2.0325
0.31171 | -34.0001 | 2.4928
0.23928 | -34.03 | 3.2105
0.18358 | -34.0282 | 3.6298
0.14091 | -34.0495 | 4.5067
0.10819 | -34.0476 | 5.6894
0.083017 | -34.0834 | 7.1576
0.063705 | -34.1256 | 9.0272
0.048923 | -34.1842 | 11.5189
0.037479 | -34.3035 | 14.7169
0.028801 | -34.4837 | 18.7456
0.022078 | -34.7842 | 24.1541
0.017023 | -35.263 | 30.6187
0.013018 | -36.0613 | 38.3194
0.010014 | -37.1843 | 46.2291

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003614 | -3.146
0.007227 | -27.2122
0.010841 | -34.6249
0.014454 | -39.365
0.018068 | -41.8756
0.021681 | -44.0125
0.025295 | -45.9474
0.028908 | -47.5372
0.032522 | -48.9967
0.036135 | -50.0053
0.039749 | -50.7797
0.043362 | -51.3854
0.046976 | -51.9369
0.050589 | -52.3179
0.054203 | -54.2752
0.057817 | -54.7942
0.06143 | -56.2746
0.065044 | -56.1204
0.068657 | -56.0767
0.072271 | -56.8991
0.075884 | -59.1205
0.079498 | -59.5175
0.083111 | -60.0181
0.086725 | -59.1973
0.090338 | -60.0096
0.093952 | -60.8667
0.097565 | -62.0153
0.10118 | -61.7483
0.10479 | -61.8321
0.10841 | -62.6964
0.11202 | -63.8572
0.11563 | -63.3854
0.11925 | -62.7748
0.12286 | -63.3262
0.12647 | -64.4011
0.13009 | -65.9148
0.1337 | -65.9226
0.13731 | -66.0986
0.14093 | -66.0014
0.14454 | -66.4049
0.14816 | -65.8574
0.15177 | -65.6914
0.15538 | -66.673
0.15899 | -66.9021
0.16261 | -67.0366
0.16622 | -67.9664
0.16984 | -67.7642
0.17345 | -67.9947
0.17706 | -68.3773
0.18068 | -68.693
0.18429 | -68.9674
0.1879 | -70.0364
0.19152 | -69.9643
0.19513 | -69.2232
0.19874 | -69.757
0.20236 | -70.5239
0.20597 | -70.5803
0.20958 | -70.7958
0.2132 | -71.0588
0.21681 | -71.5893
0.22043 | -70.933
0.22404 | -70.942
0.22765 | -71.2562
0.23127 | -72.7362
0.23488 | -72.0903
0.23849 | -71.1206
0.24211 | -71.8042
0.24572 | -72.7037
0.24933 | -72.0441
0.25295 | -73.082
0.25656 | -73.4208
0.26017 | -73.6511
0.26379 | -73.2792
0.2674 | -73.825
0.27102 | -73.731
0.27463 | -74.1698
0.27824 | -74.6701
0.28186 | -75.0875
0.28547 | -75.3831
0.28908 | -74.8736
0.2927 | -74.2926
0.29631 | -74.1792
0.29992 | -74.4054
0.30354 | -75.5384
0.30715 | -76.0511
0.31076 | -75.6219
0.31438 | -75.1784
0.31799 | -75.0616
0.3216 | -76.188
0.32522 | -76.187
0.32883 | -76.9751
0.33244 | -76.4219
0.33606 | -76.2405
0.33967 | -76.5981
0.34329 | -76.7866
0.3469 | -76.2807
0.35051 | -76.5429
0.35413 | -76.6291
0.35774 | -76.903
0.36135 | -76.2564
0.36497 | -76.8731
0.36858 | -77.6075
0.37219 | -77.7578
0.37581 | -77.6764
0.37942 | -77.7363
0.38303 | -78.381
0.38665 | -78.6297
0.39026 | -78.1635
0.39387 | -77.9158
0.39749 | -78.1118
0.4011 | -77.9505
0.40472 | -77.0494
0.40833 | -77.8169
0.41194 | -79.4654
0.41556 | -78.9961
0.41917 | -78.9107
0.42278 | -79.2739
0.4264 | -78.6617
0.43001 | -78.771
0.43362 | -79.1597
0.43724 | -79.627
0.44085 | -78.6028
0.44446 | -78.6687
0.44808 | -79.6063
0.45169 | -79.2422
0.45531 | -78.7668
0.45892 | -79.5713
0.46253 | -80.278
0.46615 | -81.0486
0.46976 | -80.447
0.47337 | -80.6703
0.47699 | -81.1598
0.4806 | -79.6847
0.48421 | -80.2556
0.48783 | -81.2078
0.49144 | -81.2703
0.49505 | -80.8189
0.49867 | -80.8841
0.50228 | -80.5277
0.50589 | -79.7884
0.50951 | -80.55
0.51312 | -81.0252
0.51674 | -81.3902
0.52035 | -81.2021
0.52396 | -80.8325
0.52758 | -80.5286
0.53119 | -80.3752
0.5348 | -80.5376
0.53842 | -81.0408
0.54203 | -81.4347
0.54564 | -81.716
0.54926 | -82.2968
0.55287 | -82.8178
0.55648 | -82.8366
0.5601 | -82.4041
0.56371 | -82.0392
0.56732 | -81.8788
0.57094 | -81.4955
0.57455 | -81.2628
0.57817 | -81.5087
0.58178 | -81.9783
0.58539 | -82.4123
0.58901 | -82.4295
0.59262 | -81.9624
0.59623 | -82.0957
0.59985 | -82.6098
0.60346 | -82.9753
0.60707 | -82.8124
0.61069 | -82.5282
0.6143 | -82.7514
0.61791 | -82.9188
0.62153 | -82.6857
0.62514 | -82.3973
0.62875 | -82.5027
0.63237 | -82.7379
0.63598 | -82.4596
0.6396 | -82.4403
0.64321 | -82.903
0.64682 | -83.5808
0.65044 | -84.372
0.65405 | -84.4911
0.65766 | -84.1552
0.66128 | -84.1073
0.66489 | -84.3444
0.6685 | -84.3162
0.67212 | -84.173
0.67573 | -84.0595
0.67934 | -83.9127
0.68296 | -83.6393
0.68657 | -83.0027
0.69019 | -82.342
0.6938 | -82.3868
0.69741 | -83.4321
0.70103 | -84.867
0.70464 | -85.6311
0.70825 | -85.3091
0.71187 | -84.6485
0.71548 | -84.186
0.71909 | -83.7195
0.72271 | -83.4025
0.72632 | -83.3352
0.72993 | -83.3527
0.73355 | -83.2703
0.73716 | -83.1933
0.74077 | -83.4092
0.74439 | -83.9174
0.748 | -84.6496
0.75162 | -85.1265
0.75523 | -84.7557
0.75884 | -84.6449
0.76245 | -84.8572
0.76607 | -85.051
0.76968 | -85.2917
0.7733 | -85.2723
0.77691 | -84.8655
0.78052 | -84.6343
0.78414 | -84.88
0.78775 | -85.3374
0.79136 | -85.5365
0.79498 | -85.1899
0.79859 | -84.8603
0.8022 | -85.1015
0.80582 | -85.675
0.80943 | -85.567
0.81305 | -84.8038
0.81666 | -84.4199
0.82027 | -84.7445
0.82388 | -85.3238
0.8275 | -85.7114
0.83111 | -85.7039
0.83473 | -85.4975
0.83834 | -85.3786
0.84195 | -85.527
0.84557 | -85.8183
0.84918 | -86.1421
0.85279 | -86.4594
0.85641 | -86.3897
0.86002 | -86.099
0.86363 | -86.1873
0.86725 | -86.5363
0.87086 | -86.3848
0.87448 | -86.3999
0.87809 | -86.6293
0.8817 | -86.7726
0.88532 | -86.6057
0.88893 | -86.3749
0.89254 | -86.2635
0.89616 | -85.9966
0.89977 | -85.736
0.90338 | -85.9441
0.907 | -86.547
0.91061 | -86.6939
0.91422 | -86.5967
0.91784 | -86.3874
0.92145 | -85.8891
0.92506 | -85.2517
0.92868 | -85.1016
0.93229 | -85.4852
0.9359 | -86.2552
0.93952 | -87.5571
0.94313 | -88.4318
0.94675 | -87.7748
0.95036 | -86.4781
0.95397 | -85.7281
0.95759 | -85.378
0.9612 | -85.1715
0.96481 | -85.1822
0.96843 | -85.4954
0.97204 | -86.0126
0.97565 | -86.4793
0.97927 | -86.8422
0.98288 | -87.4489
0.98649 | -87.9806
0.99011 | -87.7371
0.99372 | -87.3647
0.99733 | -87.2478
1.0009 | -87.4888
1.0046 | -87.7121
1.0082 | -87.7814
1.0118 | -87.6872
1.0154 | -87.2172
1.019 | -86.8808
1.0226 | -87.1723
1.0262 | -87.9753
1.0299 | -88.3464
1.0335 | -87.8927
1.0371 | -87.282
1.0407 | -87.5467
1.0443 | -87.8875
1.0479 | -88.3594
1.0515 | -88.854
1.0552 | -89.1609
1.0588 | -89.0503
1.0624 | -88.6306
1.066 | -88.3047
1.0696 | -88.2729
1.0732 | -88.341
1.0768 | -88.334
1.0804 | -88.3694
1.0841 | -88.5102
1.0877 | -88.6377
1.0913 | -88.7009
1.0949 | -88.8042
1.0985 | -88.9375
1.1021 | -88.9405
1.1057 | -88.8737
1.1094 | -88.9526
1.113 | -89.1887
1.1166 | -89.2987
1.1202 | -88.9928
1.1238 | -88.4967
1.1274 | -88.2703
1.131 | -88.4122
1.1346 | -88.7101
1.1383 | -88.8799
1.1419 | -88.9133
1.1455 | -88.9146
1.1491 | -88.7701
1.1527 | -88.5099
1.1563 | -88.441
1.1599 | -88.6351
1.1636 | -88.8304
1.1672 | -88.8648
1.1708 | -88.8607
1.1744 | -88.8128
1.178 | -88.6024
1.1816 | -88.3875
1.1852 | -88.3653
1.1889 | -88.51
1.1925 | -88.7576
1.1961 | -89.0378
1.1997 | -89.2593
1.2033 | -89.4645
1.2069 | -89.6166
1.2105 | -89.5969
1.2141 | -89.4038
1.2178 | -89.1118
1.2214 | -88.8744
1.225 | -88.7606
1.2286 | -88.7342
1.2322 | -88.8483
1.2358 | -89.0301
1.2394 | -89.1203
1.2431 | -89.1287
1.2467 | -89.0656
1.2503 | -88.8873
1.2539 | -88.6638
1.2575 | -88.6379
1.2611 | -88.9487
1.2647 | -89.3923
1.2684 | -89.6594
1.272 | -89.7028
1.2756 | -89.6273
1.2792 | -89.5428
1.2828 | -89.5446
1.2864 | -89.5918
1.29 | -89.5331
1.2936 | -89.4676
1.2973 | -89.5386
1.3009 | -89.7224
1.3045 | -89.8708
1.3081 | -89.8276
1.3117 | -89.6549
1.3153 | -89.566
1.3189 | -89.6657
1.3226 | -89.9544
1.3262 | -90.3401
1.3298 | -90.6188
1.3334 | -90.6376
1.337 | -90.6025
1.3406 | -90.6784
1.3442 | -90.7987
1.3478 | -90.9216
1.3515 | -90.8137
1.3551 | -90.3425
1.3587 | -89.8908
1.3623 | -89.8444
1.3659 | -90.214
1.3695 | -90.6339
1.3731 | -90.7168
1.3768 | -90.5765
1.3804 | -90.5552
1.384 | -90.6862
1.3876 | -90.7887
1.3912 | -90.8468
1.3948 | -90.8555
1.3984 | -90.6852
1.4021 | -90.3604
1.4057 | -90.0738
1.4093 | -90.0346
1.4129 | -90.2951
1.4165 | -90.6481
1.4201 | -90.863
1.4237 | -90.9078
1.4273 | -90.7723
1.431 | -90.4565
1.4346 | -90.1329
1.4382 | -90.0572
1.4418 | -90.3275
1.4454 | -90.734
1.449 | -91.0019
1.4526 | -91.1355
1.4563 | -91.3196
1.4599 | -91.5418
1.4635 | -91.547
1.4671 | -91.3591
1.4707 | -91.3049
1.4743 | -91.3395
1.4779 | -91.2304
1.4815 | -90.9359
1.4852 | -90.5702
1.4888 | -90.4806
1.4924 | -90.8786
1.496 | -91.4609
1.4996 | -91.6874
1.5032 | -91.4958
1.5068 | -91.1626
1.5105 | -90.8009
1.5141 | -90.5596
1.5177 | -90.6394
1.5213 | -90.9998
1.5249 | -91.3294
1.5285 | -91.3209
1.5321 | -91.1039
1.5358 | -90.9983
1.5394 | -91.042
1.543 | -91.0822
1.5466 | -91.1086
1.5502 | -91.3014
1.5538 | -91.6467
1.5574 | -91.8481
1.561 | -91.6236
1.5647 | -91.4767
1.5683 | -91.3793
1.5719 | -91.3901
1.5755 | -91.4849
1.5791 | -91.5873
1.5827 | -91.6334
1.5863 | -91.5612
1.59 | -91.4322
1.5936 | -91.3808
1.5972 | -91.4615
1.6008 | -91.6099
1.6044 | -91.732
1.608 | -91.8125
1.6116 | -91.9092
1.6152 | -92.0664
1.6189 | -92.2515
1.6225 | -92.3539
1.6261 | -92.3098
1.6297 | -92.1672
1.6333 | -92.008
1.6369 | -91.9173
1.6405 | -91.9419
1.6442 | -92.0726
1.6478 | -92.2231
1.6514 | -92.2884
1.655 | -92.2069
1.6586 | -92.0117
1.6622 | -91.811
1.6658 | -91.7088
1.6695 | -91.7653
1.6731 | -92.0007
1.6767 | -92.3745
1.6803 | -92.7376
1.6839 | -92.8381
1.6875 | -92.6153
1.6911 | -92.2484
1.6947 | -91.974
1.6984 | -91.9281
1.702 | -92.092
1.7056 | -92.3174
1.7092 | -92.414
1.7128 | -92.3769
1.7164 | -92.3782
1.72 | -92.4987
1.7237 | -92.6837
1.7273 | -92.8212
1.7309 | -92.8701
1.7345 | -92.826
1.7381 | -92.7198
1.7417 | -92.6517
1.7453 | -92.6553
1.749 | -92.6947
1.7526 | -92.7401
1.7562 | -92.7864
1.7598 | -92.787
1.7634 | -92.7059
1.767 | -92.5928
1.7706 | -92.4656
1.7742 | -92.2878
1.7779 | -92.095
1.7815 | -91.9518
1.7851 | -91.8725
1.7887 | -91.8515
1.7923 | -91.9097
1.7959 | -92.0832
1.7995 | -92.3432
1.8032 | -92.6432
1.8068 | -92.96
1.8104 | -93.2371
1.814 | -93.3506
1.8176 | -93.1904
1.8212 | -92.8653
1.8248 | -92.6312
1.8284 | -92.6053
1.8321 | -92.7483
1.8357 | -92.9238
1.8393 | -92.9651
1.8429 | -92.7797
1.8465 | -92.5018
1.8501 | -92.349
1.8537 | -92.3724
1.8574 | -92.4922
1.861 | -92.6216
1.8646 | -92.6825
1.8682 | -92.601
1.8718 | -92.3783
1.8754 | -92.1663
1.879 | -92.1236
1.8827 | -92.3277
1.8863 | -92.7077
1.8899 | -93.0642
1.8935 | -93.1976
1.8971 | -93.0677
1.9007 | -92.8281
1.9043 | -92.6878
1.9079 | -92.7045
1.9116 | -92.8754
1.9152 | -93.1655
1.9188 | -93.4781
1.9224 | -93.6359
1.926 | -93.5796
1.9296 | -93.4309
1.9332 | -93.3406
1.9369 | -93.3086
1.9405 | -93.2756
1.9441 | -93.1869
1.9477 | -93.0486
1.9513 | -92.9102
1.9549 | -92.8237
1.9585 | -92.8456
1.9621 | -93.0025
1.9658 | -93.1984
1.9694 | -93.3173
1.973 | -93.3146
1.9766 | -93.2538
1.9802 | -93.2371
1.9838 | -93.3284
1.9874 | -93.527
1.9911 | -93.7423
1.9947 | -93.84
1.9983 | -93.7749
2.0019 | -93.6169
2.0055 | -93.5106
2.0091 | -93.5456
2.0127 | -93.7068
2.0164 | -93.9213
2.02 | -94.0836
2.0236 | -94.0971
2.0272 | -93.9819
2.0308 | -93.8014
2.0344 | -93.617
2.038 | -93.4636
2.0416 | -93.411
2.0453 | -93.4801
2.0489 | -93.622
2.0525 | -93.7374
2.0561 | -93.7406
2.0597 | -93.6416
2.0633 | -93.562
2.0669 | -93.5781
2.0706 | -93.6767
2.0742 | -93.7309
2.0778 | -93.6351
2.0814 | -93.4077
2.085 | -93.3388
2.0886 | -93.3523
2.0922 | -93.5015
2.0958 | -93.7533
2.0995 | -94.0245
2.1031 | -94.234
2.1067 | -94.3297
2.1103 | -94.3494
2.1139 | -94.3661
2.1175 | -94.4472
2.1211 | -94.606
2.1248 | -94.8155
2.1284 | -94.9874
2.132 | -95.0175
2.1356 | -94.8588
2.1392 | -94.5747
2.1428 | -94.3033
2.1464 | -94.1575
2.1501 | -94.169
2.1537 | -94.3008
2.1573 | -94.4523
2.1609 | -94.5527
2.1645 | -94.5845
2.1681 | -94.5869
2.1717 | -94.6063
2.1753 | -94.6607
2.179 | -94.7064
2.1826 | -94.6826
2.1862 | -94.571
2.1898 | -94.4141
2.1934 | -94.2789
2.197 | -94.2
2.2006 | -94.1793
2.2043 | -94.2133
2.2079 | -94.2854
2.2115 | -94.3827
2.2151 | -94.5037
2.2187 | -94.636
2.2223 | -94.764
2.2259 | -94.8658
2.2296 | -94.9171
2.2332 | -94.8977
2.2368 | -94.8085
2.2404 | -94.6693
2.244 | -94.5545
2.2476 | -94.5074
2.2512 | -94.5197
2.2548 | -94.5509
2.2585 | -94.5966
2.2621 | -94.677
2.2657 | -94.8116
2.2693 | -94.9588
2.2729 | -95.0427
2.2765 | -95.0232
2.2801 | -94.9331
2.2838 | -94.813
2.2874 | -94.7174
2.291 | -94.6799
2.2946 | -94.6794
2.2982 | -94.6786
2.3018 | -94.6448
2.3054 | -94.5829
2.309 | -94.5147
2.3127 | -94.4716
2.3163 | -94.4706
2.3199 | -94.5144
2.3235 | -94.6214
2.3271 | -94.7834
2.3307 | -94.9722
2.3343 | -95.1456
2.338 | -95.2539
2.3416 | -95.2944
2.3452 | -95.2933
2.3488 | -95.2907
2.3524 | -95.2865
2.356 | -95.2827
2.3596 | -95.2786
2.3632 | -95.2549
2.3669 | -95.2047
2.3705 | -95.1266
2.3741 | -95.0625
2.3777 | -95.0325
2.3813 | -95.0442
2.3849 | -95.0896
2.3885 | -95.1485
2.3922 | -95.1761
2.3958 | -95.1352
2.3994 | -95.0592
2.403 | -95.0011
2.4066 | -94.998
2.4102 | -95.0534
2.4138 | -95.1449
2.4175 | -95.2434
2.4211 | -95.3323
2.4247 | -95.386
2.4283 | -95.3587
2.4319 | -95.2455
2.4355 | -95.0701
2.4391 | -94.9105
2.4427 | -94.8349
2.4464 | -94.8667
2.45 | -94.9878
2.4536 | -95.1617
2.4572 | -95.357
2.4608 | -95.5054
2.4644 | -95.5686
2.468 | -95.5571
2.4717 | -95.4709
2.4753 | -95.3365
2.4789 | -95.175
2.4825 | -95.0145
2.4861 | -94.8941
2.4897 | -94.8167
2.4933 | -94.7723
2.497 | -94.7658
2.5006 | -94.8036
2.5042 | -94.8852
2.5078 | -94.9833
2.5114 | -95.078
2.515 | -95.1382
2.5186 | -95.1385
2.5222 | -95.12
2.5259 | -95.1156
2.5295 | -95.1608
2.5331 | -95.2179
2.5367 | -95.2336
2.5403 | -95.1999
2.5439 | -95.1463
2.5475 | -95.1011
2.5512 | -95.0743
2.5548 | -95.0534
2.5584 | -95.0301
2.562 | -95.0093
2.5656 | -94.9767
2.5692 | -94.9106
2.5728 | -94.8251
2.5764 | -94.7482
2.5801 | -94.7143
2.5837 | -94.7514
2.5873 | -94.8892
2.5909 | -95.109
2.5945 | -95.3676
2.5981 | -95.5822
2.6017 | -95.6608
2.6054 | -95.531
2.609 | -95.4676
2.6126 | -95.4134
2.6162 | -95.403
2.6198 | -95.4393
2.6234 | -95.5217
2.627 | -95.6085
2.6307 | -95.6446
2.6343 | -95.6221
2.6379 | -95.5477
2.6415 | -95.4629
2.6451 | -95.4166
2.6487 | -95.4269
2.6523 | -95.4715
2.6559 | -95.5149
2.6596 | -95.5408
2.6632 | -95.5573
2.6668 | -95.5943
2.6704 | -95.6666
2.674 | -95.7765
2.6776 | -95.89
2.6812 | -95.939
2.6849 | -95.8784
2.6885 | -95.7237
2.6921 | -95.5336
2.6957 | -95.392
2.6993 | -95.3467
2.7029 | -95.4146
2.7065 | -95.5678
2.7102 | -95.7646
2.7138 | -95.9641
2.7174 | -96.1251
2.721 | -96.2321
2.7246 | -96.2821
2.7282 | -96.2928
2.7318 | -96.2816
2.7354 | -96.2465
2.7391 | -96.2022
2.7427 | -96.141
2.7463 | -96.0687
2.7499 | -95.9881
2.7535 | -95.9342
2.7571 | -95.9187
2.7607 | -95.9488
2.7644 | -96.0326
2.768 | -96.1485
2.7716 | -96.2679
2.7752 | -96.3515
2.7788 | -96.3804
2.7824 | -96.3499
2.786 | -96.265
2.7896 | -96.1598
2.7933 | -96.0631
2.7969 | -96.0108
2.8005 | -96.0302
2.8041 | -96.1298
2.8077 | -96.2729
2.8113 | -96.4045
2.8149 | -96.4946
2.8186 | -96.5372
2.8222 | -96.5498
2.8258 | -96.5377
2.8294 | -96.5045
2.833 | -96.4534
2.8366 | -96.3972
2.8402 | -96.3537
2.8439 | -96.3426
2.8475 | -96.3644
2.8511 | -96.4182
2.8547 | -96.4915
2.8583 | -96.5788
2.8619 | -96.6578
2.8655 | -96.6905
2.8691 | -96.6658
2.8728 | -96.5905
2.8764 | -96.4982
2.88 | -96.4302
2.8836 | -96.4141
2.8872 | -96.4661
2.8908 | -96.56
2.8944 | -96.6686
2.8981 | -96.7829
2.9017 | -96.8927
2.9053 | -96.9763
2.9089 | -97.0451
2.9125 | -97.0983
2.9161 | -97.1015
2.9197 | -97.0554
2.9233 | -96.9595
2.927 | -96.806
2.9306 | -96.614
2.9342 | -96.4127
2.9378 | -96.2571
2.9414 | -96.1519
2.945 | -96.109
2.9486 | -96.1218
2.9523 | -96.1856
2.9559 | -96.2844
2.9595 | -96.41
2.9631 | -96.5539
2.9667 | -96.676
2.9703 | -96.746
2.9739 | -96.7632
2.9776 | -96.7169
2.9812 | -96.6335
2.9848 | -96.5486
2.9884 | -96.5116
2.992 | -96.5317
2.9956 | -96.6103
2.9992 | -96.7179
3.0028 | -96.8167
3.0065 | -96.8945
3.0101 | -96.9531
3.0137 | -97.0056
3.0173 | -97.0576
3.0209 | -97.1211
3.0245 | -97.1988
3.0281 | -97.2649
3.0318 | -97.322
3.0354 | -97.3576
3.039 | -97.3427
3.0426 | -97.2718
3.0462 | -97.1454
3.0498 | -96.98
3.0534 | -96.8004
3.057 | -96.642
3.0607 | -96.5292
3.0643 | -96.4791
3.0679 | -96.4861
3.0715 | -96.5524
3.0751 | -96.6652
3.0787 | -96.7916
3.0823 | -96.934
3.086 | -97.0765
3.0896 | -97.1878
3.0932 | -97.2599
3.0968 | -97.2776
3.1004 | -97.24
3.104 | -97.1639
3.1076 | -97.068
3.1113 | -96.9842
3.1149 | -96.9298
3.1185 | -96.9078
3.1221 | -96.9097
3.1257 | -96.9368
3.1293 | -96.9599
3.1329 | -96.9712
3.1365 | -96.982
3.1402 | -96.9926
3.1438 | -96.9733
3.1474 | -96.9367
3.151 | -96.8915
3.1546 | -96.8421
3.1582 | -96.7932
3.1618 | -96.7504
3.1655 | -96.7218
3.1691 | -96.6951
3.1727 | -96.6682
3.1763 | -96.6449
3.1799 | -96.6361
3.1835 | -96.658
3.1871 | -96.7293
3.1907 | -96.8313
3.1944 | -96.9504
3.198 | -97.0889
3.2016 | -97.2104
3.2052 | -97.3041
3.2088 | -97.3759
3.2124 | -97.4127
3.216 | -97.4174
3.2197 | -97.4117
3.2233 | -97.4093
3.2269 | -97.4221
3.2305 | -97.4378
3.2341 | -97.4612
3.2377 | -97.4684
3.2413 | -97.4482
3.245 | -97.4191
3.2486 | -97.3944
3.2522 | -97.365
3.2558 | -97.362
3.2594 | -97.3859
3.263 | -97.4385
3.2666 | -97.5227
3.2702 | -97.6149
3.2739 | -97.6935
3.2775 | -97.7494
3.2811 | -97.7591
3.2847 | -97.7264
3.2883 | -97.6584
3.2919 | -97.585
3.2955 | -97.5197
3.2992 | -97.4643
3.3028 | -97.4169
3.3064 | -97.3756
3.31 | -97.3273
3.3136 | -97.277
3.3172 | -97.2311
3.3208 | -97.1985
3.3245 | -97.188
3.3281 | -97.1992
3.3317 | -97.2212
3.3353 | -97.2506
3.3389 | -97.2814
3.3425 | -97.3165
3.3461 | -97.3594
3.3497 | -97.4191
3.3534 | -97.4877
3.357 | -97.5525
3.3606 | -97.6083
3.3642 | -97.6571
3.3678 | -97.7161
3.3714 | -97.7918
3.375 | -97.8817
3.3787 | -97.9664
3.3823 | -98.0174
3.3859 | -98.0145
3.3895 | -97.9664
3.3931 | -97.8857
3.3967 | -97.7977
3.4003 | -97.7166
3.4039 | -97.6498
3.4076 | -97.6223
3.4112 | -97.6369
3.4148 | -97.6933
3.4184 | -97.778
3.422 | -97.8768
3.4256 | -97.9776
3.4292 | -98.0613
3.4329 | -98.122
3.4365 | -98.1671
3.4401 | -98.1948
3.4437 | -98.1958
3.4473 | -98.1864
3.4509 | -98.1771
3.4545 | -98.1879
3.4582 | -98.2151
3.4618 | -98.2506
3.4654 | -98.2941
3.469 | -98.3299
3.4726 | -98.3426
3.4762 | -98.335
3.4798 | -98.3173
3.4834 | -98.2934
3.4871 | -98.2549
3.4907 | -98.1878
3.4943 | -98.1037
3.4979 | -98.0284
3.5015 | -97.9885
3.5051 | -97.9913
3.5087 | -98.0378
3.5124 | -98.1131
3.516 | -98.1819
3.5196 | -98.2201
3.5232 | -98.2157
3.5268 | -98.1691
3.5304 | -98.0837
3.534 | -97.9549
3.5376 | -97.8029
3.5413 | -97.6484
3.5449 | -97.5179
3.5485 | -97.4241
3.5521 | -97.3844
3.5557 | -97.4082
3.5593 | -97.506
3.5629 | -97.6682
3.5666 | -97.8651
3.5702 | -98.0683
3.5738 | -98.239
3.5774 | -98.3395
3.581 | -98.3512
3.5846 | -98.2839
3.5882 | -98.1746
3.5919 | -98.0681
3.5955 | -98.0011
3.5991 | -97.9924
3.6027 | -98.0482
3.6063 | -98.1612
3.6099 | -98.3135
3.6135 | -98.4745
3.6171 | -98.6245
3.6208 | -98.7343
3.6244 | -98.7751
3.628 | -98.7574
3.6316 | -98.7086
3.6352 | -98.6649
3.6388 | -98.645
3.6424 | -98.668
3.6461 | -98.7466
3.6497 | -98.7909
3.6533 | -98.8377
3.6569 | -98.8806
3.6605 | -98.9025
3.6641 | -98.8939
3.6677 | -98.8551
3.6713 | -98.7851
3.675 | -98.6852
3.6786 | -98.5883
3.6822 | -98.5144
3.6858 | -98.4681
3.6894 | -98.4532
3.693 | -98.459
3.6966 | -98.4782
3.7003 | -98.5019
3.7039 | -98.5318
3.7075 | -98.569
3.7111 | -98.6158
3.7147 | -98.6763
3.7183 | -98.7475
3.7219 | -98.8318
3.7256 | -98.9175
3.7292 | -98.9935
3.7328 | -99.0436
3.7364 | -99.0584
3.74 | -99.0107
3.7436 | -98.9088
3.7472 | -98.7859
3.7508 | -98.6546
3.7545 | -98.5407
3.7581 | -98.4651
3.7617 | -98.4328
3.7653 | -98.4435
3.7689 | -98.4957
3.7725 | -98.5786
3.7761 | -98.6727
3.7798 | -98.7744
3.7834 | -98.8621
3.787 | -98.9244
3.7906 | -98.9554
3.7942 | -98.9607
3.7978 | -98.9424
3.8014 | -98.9321
3.8051 | -98.9249
3.8087 | -98.9156
3.8123 | -98.9152
3.8159 | -98.908
3.8195 | -98.8805
3.8231 | -98.8334
3.8267 | -98.7787
3.8303 | -98.7139
3.834 | -98.6448
3.8376 | -98.5886
3.8412 | -98.536
3.8448 | -98.4971
3.8484 | -98.4952
3.852 | -98.5328
3.8556 | -98.6053
3.8593 | -98.7128
3.8629 | -98.833
3.8665 | -98.9595
3.8701 | -99.0743
3.8737 | -99.1482
3.8773 | -99.1739
3.8809 | -99.1583
3.8845 | -99.1059
3.8882 | -99.0288
3.8918 | -98.9519
3.8954 | -98.8784
3.899 | -98.8117
3.9026 | -98.7566
3.9062 | -98.7043
3.9098 | -98.6536
3.9135 | -98.6167
3.9171 | -98.5996
3.9207 | -98.5932
3.9243 | -98.613
3.9279 | -98.6495
3.9315 | -98.6915
3.9351 | -98.7352
3.9388 | -98.7631
3.9424 | -98.7593
3.946 | -98.7325
3.9496 | -98.6877
3.9532 | -98.643
3.9568 | -98.6171
3.9604 | -98.6191
3.964 | -98.6599
3.9677 | -98.7365
3.9713 | -98.8345
3.9749 | -98.9477
3.9785 | -99.0603
3.9821 | -99.1588
3.9857 | -99.2211
3.9893 | -99.267
3.993 | -99.3075
3.9966 | -99.3415
4.0002 | -99.3753
4.0038 | -99.4111
4.0074 | -99.4391
4.011 | -99.4687
4.0146 | -99.4933
4.0182 | -99.5152
4.0219 | -99.5317
4.0255 | -99.5396
4.0291 | -99.5406
4.0327 | -99.5367
4.0363 | -99.52
4.0399 | -99.4942
4.0435 | -99.4617
4.0472 | -99.4233
4.0508 | -99.3675
4.0544 | -99.3025
4.058 | -99.2451
4.0616 | -99.1956
4.0652 | -99.167
4.0688 | -99.1496
4.0725 | -99.1361
4.0761 | -99.1384
4.0797 | -99.1496
4.0833 | -99.1644
4.0869 | -99.1821
4.0905 | -99.1953
4.0941 | -99.1903
4.0977 | -99.1766
4.1014 | -99.1504
4.105 | -99.1085
4.1086 | -99.0528
4.1122 | -98.987
4.1158 | -98.9201
4.1194 | -98.8642
4.123 | -98.8285
4.1267 | -98.8319
4.1303 | -98.8794
4.1339 | -98.9573
4.1375 | -99.0451
4.1411 | -99.1321
4.1447 | -99.2019
4.1483 | -99.2413
4.1519 | -99.2554
4.1556 | -99.2584
4.1592 | -99.2517
4.1628 | -99.2264
4.1664 | -99.1878
4.17 | -99.145
4.1736 | -99.1033
4.1772 | -99.0688
4.1809 | -99.0451
4.1845 | -99.0364
4.1881 | -99.055
4.1917 | -99.0907
4.1953 | -99.1534
4.1989 | -99.2211
4.2025 | -99.2846
4.2062 | -99.3344
4.2098 | -99.3617
4.2134 | -99.3662
4.217 | -99.353
4.2206 | -99.3166
4.2242 | -99.2604
4.2278 | -99.209
4.2314 | -99.1605
4.2351 | -99.1105
4.2387 | -99.0815
4.2423 | -99.0705
4.2459 | -99.0804
4.2495 | -99.1195
4.2531 | -99.1911
4.2567 | -99.2877
4.2604 | -99.4053
4.264 | -99.5303
4.2676 | -99.6571
4.2712 | -99.7501
4.2748 | -99.8005
4.2784 | -99.8116
4.282 | -99.7871
4.2857 | -99.7364
4.2893 | -99.6896
4.2929 | -99.6562
4.2965 | -99.6284
4.3001 | -99.6032
4.3037 | -99.6
4.3073 | -99.6015
4.3109 | -99.6065
4.3146 | -99.609
4.3182 | -99.6041
4.3218 | -99.5927
4.3254 | -99.5801
4.329 | -99.5666
4.3326 | -99.5697
4.3362 | -99.5876
4.3399 | -99.6367
4.3435 | -99.712
4.3471 | -99.7986
4.3507 | -99.9002
4.3543 | -100.0031
4.3579 | -100.0959
4.3615 | -100.1745
4.3651 | -100.2382
4.3688 | -100.2847
4.3724 | -100.294
4.376 | -100.2779
4.3796 | -100.2391
4.3832 | -100.1795
4.3868 | -100.1092
4.3904 | -100.0414
4.3941 | -99.9791
4.3977 | -99.925
4.4013 | -99.8819
4.4049 | -99.8563
4.4085 | -99.8475
4.4121 | -99.8441
4.4157 | -99.8546
4.4193 | -99.8635
4.423 | -99.881
4.4266 | -99.908
4.4302 | -99.9484
4.4338 | -100.0028
4.4374 | -100.0615
4.441 | -100.1231
4.4446 | -100.1847
4.4483 | -100.2349
4.4519 | -100.2658
4.4555 | -100.2765
4.4591 | -100.2727
4.4627 | -100.2695
4.4663 | -100.2754
4.4699 | -100.2877
4.4736 | -100.3139
4.4772 | -100.3438
4.4808 | -100.3917
4.4844 | -100.4275
4.488 | -100.4509
4.4916 | -100.4545
4.4952 | -100.431
4.4988 | -100.3801
4.5025 | -100.3103
4.5061 | -100.2307
4.5097 | -100.1402
4.5133 | -100.045
4.5169 | -99.9509
4.5205 | -99.8692
4.5241 | -99.811
4.5278 | -99.7688
4.5314 | -99.7355
4.535 | -99.7104
4.5386 | -99.6977
4.5422 | -99.6984
4.5458 | -99.711
4.5494 | -99.7387
4.5531 | -99.7869
4.5567 | -99.843
4.5603 | -99.9124
4.5639 | -99.9837
4.5675 | -100.0504
4.5711 | -100.1026
4.5747 | -100.1442
4.5783 | -100.1754
4.582 | -100.1952
4.5856 | -100.2068
4.5892 | -100.2125
4.5928 | -100.2092
4.5964 | -100.2073
4.6 | -100.2145
4.6036 | -100.2257
4.6073 | -100.2389
4.6109 | -100.2538
4.6145 | -100.2785
4.6181 | -100.31
4.6217 | -100.3603
4.6253 | -100.4282
4.6289 | -100.4993
4.6325 | -100.5724
4.6362 | -100.6419
4.6398 | -100.6966
4.6434 | -100.7276
4.647 | -100.7319
4.6506 | -100.7157
4.6542 | -100.6874
4.6578 | -100.6501
4.6615 | -100.6163
4.6651 | -100.6038
4.6687 | -100.6098
4.6723 | -100.6361
4.6759 | -100.6772
4.6795 | -100.7294
4.6831 | -100.7776
4.6868 | -100.8214
4.6904 | -100.8338
4.694 | -100.8522
4.6976 | -100.8587
4.7012 | -100.8584
4.7048 | -100.8547
4.7084 | -100.8423
4.712 | -100.8257
4.7157 | -100.8147
4.7193 | -100.8147
4.7229 | -100.8335
4.7265 | -100.872
4.7301 | -100.9222
4.7337 | -100.978
4.7373 | -101.0291
4.741 | -101.0728
4.7446 | -101.1077
4.7482 | -101.1302
4.7518 | -101.1415
4.7554 | -101.1395
4.759 | -101.1169
4.7626 | -101.0837
4.7663 | -101.0506
4.7699 | -101.0191
4.7735 | -100.9894
4.7771 | -100.9672
4.7807 | -100.9561
4.7843 | -100.9553
4.7879 | -100.9647
4.7915 | -100.9759
4.7952 | -100.9825
4.7988 | -100.9747
4.8024 | -100.958
4.806 | -100.9439
4.8096 | -100.9186
4.8132 | -100.8832
4.8168 | -100.8493
4.8205 | -100.807
4.8241 | -100.7621
4.8277 | -100.7203
4.8313 | -100.6943
4.8349 | -100.6739
4.8385 | -100.6535
4.8421 | -100.6384
4.8457 | -100.6325
4.8494 | -100.635
4.853 | -100.6423
4.8566 | -100.6534
4.8602 | -100.6646
4.8638 | -100.6797
4.8674 | -100.6965
4.871 | -100.704
4.8747 | -100.705
4.8783 | -100.7036
4.8819 | -100.6932
4.8855 | -100.6858
4.8891 | -100.6776
4.8927 | -100.675
4.8963 | -100.6886
4.8999 | -100.7054
4.9036 | -100.7221
4.9072 | -100.738
4.9108 | -100.7526
4.9144 | -100.7546
4.918 | -100.739
4.9216 | -100.7211
4.9252 | -100.7018
4.9289 | -100.6823
4.9325 | -100.6757
4.9361 | -100.6838
4.9397 | -100.7041
4.9433 | -100.7439
4.9469 | -100.8106
4.9505 | -100.8836
4.9542 | -100.9448
4.9578 | -100.9995
4.9614 | -101.0402
4.965 | -101.0599
4.9686 | -101.0558
4.9722 | -101.0338
4.9758 | -101.0051
4.9794 | -100.9736
4.9831 | -100.9498
4.9867 | -100.9491
4.9903 | -100.9798
4.9939 | -101.0421
4.9975 | -101.1208
5.0011 | -101.2246
5.0047 | -101.3416
5.0084 | -101.4536
5.012 | -101.553
5.0156 | -101.6148
5.0192 | -101.6309
5.0228 | -101.6052
5.0264 | -101.5466
5.03 | -101.4699
5.0337 | -101.3745
5.0373 | -101.2663
5.0409 | -101.159
5.0445 | -101.07
5.0481 | -101.003
5.0517 | -100.9508
5.0553 | -100.9129
5.0589 | -100.887
5.0626 | -100.8664
5.0662 | -100.8575
5.0698 | -100.8615
5.0734 | -100.8763
5.077 | -100.8922
5.0806 | -100.9133
5.0842 | -100.9357
5.0879 | -100.9555
5.0915 | -100.9847
5.0951 | -101.0171
5.0987 | -101.0468
5.1023 | -101.0749
5.1059 | -101.0944
5.1095 | -101.1131
5.1131 | -101.1311
5.1168 | -101.1453
5.1204 | -101.1464
5.124 | -101.144
5.1276 | -101.1417
5.1312 | -101.1378
5.1348 | -101.1352
5.1384 | -101.1321
5.1421 | -101.126
5.1457 | -101.1188
5.1493 | -101.1174
5.1529 | -101.1176
5.1565 | -101.1122
5.1601 | -101.1067
5.1637 | -101.1048
5.1674 | -101.1054
5.171 | -101.1122
5.1746 | -101.1213
5.1782 | -101.1322
5.1818 | -101.1462
5.1854 | -101.1545
5.189 | -101.1599
5.1926 | -101.1558
5.1963 | -101.1421
5.1999 | -101.1241
5.2035 | -101.0986
5.2071 | -101.0714
5.2107 | -101.0453
5.2143 | -101.0261
5.2179 | -101.0145
5.2216 | -101.0101
5.2252 | -101.0167
5.2288 | -101.0315
5.2324 | -101.0667
5.236 | -101.1144
5.2396 | -101.1693
5.2432 | -101.2204
5.2468 | -101.2643
5.2505 | -101.3097
5.2541 | -101.3515
5.2577 | -101.3816
5.2613 | -101.3996
5.2649 | -101.4052
5.2685 | -101.3982
5.2721 | -101.3853
5.2758 | -101.3631
5.2794 | -101.3263
5.283 | -101.2838
5.2866 | -101.2364
5.2902 | -101.1885
5.2938 | -101.1483
5.2974 | -101.1163
5.3011 | -101.1005
5.3047 | -101.0981
5.3083 | -101.1078
5.3119 | -101.1317
5.3155 | -101.1552
5.3191 | -101.1861
5.3227 | -101.2211
5.3263 | -101.2652
5.33 | -101.3178
5.3336 | -101.3701
5.3372 | -101.4302
5.3408 | -101.5021
5.3444 | -101.5729
5.348 | -101.6418
5.3516 | -101.7
5.3553 | -101.7521
5.3589 | -101.7964
5.3625 | -101.83
5.3661 | -101.8498
5.3697 | -101.8458
5.3733 | -101.8252
5.3769 | -101.785
5.3806 | -101.7291
5.3842 | -101.6547
5.3878 | -101.5691
5.3914 | -101.4759
5.395 | -101.3826
5.3986 | -101.3019
5.4022 | -101.2344
5.4058 | -101.1831
5.4095 | -101.1521
5.4131 | -101.1352
5.4167 | -101.1331
5.4203 | -101.1539
5.4239 | -101.1864
5.4275 | -101.2291
5.4311 | -101.2742
5.4348 | -101.322
5.4384 | -101.3632
5.442 | -101.3886
5.4456 | -101.3984
5.4492 | -101.3993
5.4528 | -101.3879
5.4564 | -101.3596
5.46 | -101.326
5.4637 | -101.2914
5.4673 | -101.2643
5.4709 | -101.2446
5.4745 | -101.2308
5.4781 | -101.2204
5.4817 | -101.2136
5.4853 | -101.2137
5.489 | -101.2194
5.4926 | -101.2318
5.4962 | -101.2525
5.4998 | -101.2778
5.5034 | -101.303
5.507 | -101.3378
5.5106 | -101.3779
5.5143 | -101.4192
5.5179 | -101.458
5.5215 | -101.4963
5.5251 | -101.5365
5.5287 | -101.5793
5.5323 | -101.6056
5.5359 | -101.6319
5.5395 | -101.6556
5.5432 | -101.6623
5.5468 | -101.6549
5.5504 | -101.6412
5.554 | -101.626
5.5576 | -101.6128
5.5612 | -101.5995
5.5648 | -101.5976
5.5685 | -101.605
5.5721 | -101.6272
5.5757 | -101.6662
5.5793 | -101.7168
5.5829 | -101.775
5.5865 | -101.8347
5.5901 | -101.8924
5.5937 | -101.948
5.5974 | -101.9981
5.601 | -102.0392
5.6046 | -102.0643
5.6082 | -102.0763
5.6118 | -102.0704
5.6154 | -102.0539
5.619 | -102.0224
5.6227 | -101.9722
5.6263 | -101.9136
5.6299 | -101.8442
5.6335 | -101.7509
5.6371 | -101.6504
5.6407 | -101.5462
5.6443 | -101.4468
5.648 | -101.3471
5.6516 | -101.2587
5.6552 | -101.1793
5.6588 | -101.1157
5.6624 | -101.077
5.666 | -101.0622
5.6696 | -101.0724
5.6732 | -101.1022
5.6769 | -101.1519
5.6805 | -101.2134
5.6841 | -101.2859
5.6877 | -101.3629
5.6913 | -101.4396
5.6949 | -101.5161
5.6985 | -101.5907
5.7022 | -101.6469
5.7058 | -101.6897
5.7094 | -101.7056
5.713 | -101.7001
5.7166 | -101.677
5.7202 | -101.6261
5.7238 | -101.5701
5.7274 | -101.5073
5.7311 | -101.4563
5.7347 | -101.4079
5.7383 | -101.368
5.7419 | -101.3361
5.7455 | -101.3132
5.7491 | -101.2932
5.7527 | -101.2819
5.7564 | -101.2738
5.76 | -101.2684
5.7636 | -101.2733
5.7672 | -101.2854
5.7708 | -101.2976
5.7744 | -101.3179
5.778 | -101.3423
5.7817 | -101.3694
5.7853 | -101.3978
5.7889 | -101.4336
5.7925 | -101.4702
5.7961 | -101.5107
5.7997 | -101.5517
5.8033 | -101.602
5.8069 | -101.6505
5.8106 | -101.6848
5.8142 | -101.7172
5.8178 | -101.7502
5.8214 | -101.7797
5.825 | -101.8026
5.8286 | -101.8288
5.8322 | -101.8574
5.8359 | -101.8851
5.8395 | -101.9212
5.8431 | -101.9542
5.8467 | -101.9871
5.8503 | -102.0234
5.8539 | -102.0499
5.8575 | -102.0698
5.8612 | -102.0861
5.8648 | -102.0928
5.8684 | -102.0896
5.872 | -102.0834
5.8756 | -102.0802
5.8792 | -102.0744
5.8828 | -102.0751
5.8864 | -102.0782
5.8901 | -102.0813
5.8937 | -102.0813
5.8973 | -102.0863
5.9009 | -102.0904
5.9045 | -102.0903
5.9081 | -102.0829
5.9117 | -102.0715
5.9154 | -102.0548
5.919 | -102.0294
5.9226 | -102.0003
5.9262 | -101.968
5.9298 | -101.9375
5.9334 | -101.9172
5.937 | -101.9026
5.9406 | -101.8943
5.9443 | -101.8913
5.9479 | -101.8933
5.9515 | -101.8937
5.9551 | -101.8881
5.9587 | -101.8819
5.9623 | -101.8657
5.9659 | -101.8436
5.9696 | -101.8168
5.9732 | -101.7738
5.9768 | -101.7269
5.9804 | -101.6842
5.984 | -101.6441
5.9876 | -101.6109
5.9912 | -101.5953
5.9949 | -101.5909
5.9985 | -101.5973
6.0021 | -101.6244
6.0057 | -101.6682
6.0093 | -101.7237
6.0129 | -101.7891
6.0165 | -101.8609
6.0201 | -101.9281
6.0238 | -101.9902
6.0274 | -102.0566
6.031 | -102.1201
6.0346 | -102.1843
6.0382 | -102.2437
6.0418 | -102.3033
6.0454 | -102.3614
6.0491 | -102.4126
6.0527 | -102.4629
6.0563 | -102.5049
6.0599 | -102.5447
6.0635 | -102.5894
6.0671 | -102.6271
6.0707 | -102.6547
6.0743 | -102.6746
6.078 | -102.6882
6.0816 | -102.6946
6.0852 | -102.6937
6.0888 | -102.6863
6.0924 | -102.6675
6.096 | -102.6451
6.0996 | -102.6209
6.1033 | -102.5962
6.1069 | -102.5723
6.1105 | -102.5566
6.1141 | -102.5445
6.1177 | -102.5278
6.1213 | -102.51
6.1249 | -102.4994
6.1286 | -102.4937
6.1322 | -102.4842
6.1358 | -102.4719
6.1394 | -102.453
6.143 | -102.4229
6.1466 | -102.3924
6.1502 | -102.3544
6.1538 | -102.3089
6.1575 | -102.2631
6.1611 | -102.2151
6.1647 | -102.1719
6.1683 | -102.128
6.1719 | -102.0852
6.1755 | -102.0514
6.1791 | -102.0256
6.1828 | -102.0038
6.1864 | -101.9866
6.19 | -101.972
6.1936 | -101.9624
6.1972 | -101.9567
6.2008 | -101.9602
6.2044 | -101.974
6.208 | -101.9961
6.2117 | -102.0325
6.2153 | -102.0699
6.2189 | -102.1112
6.2225 | -102.1563
6.2261 | -102.2035
6.2297 | -102.2552
6.2333 | -102.3044
6.237 | -102.347
6.2406 | -102.3824
6.2442 | -102.4117
6.2478 | -102.4379
6.2514 | -102.4633
6.255 | -102.4821
6.2586 | -102.5034
6.2623 | -102.5259
6.2659 | -102.5453
6.2695 | -102.5564
6.2731 | -102.5564
6.2767 | -102.5499
6.2803 | -102.5386
6.2839 | -102.521
6.2875 | -102.4947
6.2912 | -102.4599
6.2948 | -102.4186
6.2984 | -102.3791
6.302 | -102.34
6.3056 | -102.3033
6.3092 | -102.2776
6.3128 | -102.2569
6.3165 | -102.2363
6.3201 | -102.2132
6.3237 | -102.1917
6.3273 | -102.1764
6.3309 | -102.1613
6.3345 | -102.1464
6.3381 | -102.1361
6.3418 | -102.1264
6.3454 | -102.1187
6.349 | -102.1107
6.3526 | -102.106
6.3562 | -102.1016
6.3598 | -102.0938
6.3634 | -102.0904
6.367 | -102.0942
6.3707 | -102.1073
6.3743 | -102.1239
6.3779 | -102.1423
6.3815 | -102.1702
6.3851 | -102.2042
6.3887 | -102.2445
6.3923 | -102.2946
6.396 | -102.3503
6.3996 | -102.4153
6.4032 | -102.4844
6.4068 | -102.5567
6.4104 | -102.6364
6.414 | -102.7236
6.4176 | -102.8139
6.4212 | -102.8903
6.4249 | -102.9548
6.4285 | -103.0112
6.4321 | -103.0532
6.4357 | -103.0743
6.4393 | -103.0734
6.4429 | -103.0526
6.4465 | -103.0155
6.4502 | -102.9611
6.4538 | -102.897
6.4574 | -102.8268
6.461 | -102.753
6.4646 | -102.6808
6.4682 | -102.615
6.4718 | -102.5567
6.4755 | -102.5122
6.4791 | -102.4831
6.4827 | -102.4699
6.4863 | -102.4688
6.4899 | -102.4729
6.4935 | -102.4871
6.4971 | -102.5158
6.5007 | -102.5452
6.5044 | -102.5835
6.508 | -102.6319
6.5116 | -102.6797
6.5152 | -102.7282
6.5188 | -102.7799
6.5224 | -102.8285
6.526 | -102.871
6.5297 | -102.9074
6.5333 | -102.9299
6.5369 | -102.9437
6.5405 | -102.9447
6.5441 | -102.9308
6.5477 | -102.912
6.5513 | -102.8882
6.5549 | -102.8596
6.5586 | -102.8173
6.5622 | -102.774
6.5658 | -102.7266
6.5694 | -102.6836
6.573 | -102.6461
6.5766 | -102.6118
6.5802 | -102.5877
6.5839 | -102.5757
6.5875 | -102.5784
6.5911 | -102.5907
6.5947 | -102.6121
6.5983 | -102.6493
6.6019 | -102.6938
6.6055 | -102.743
6.6092 | -102.8002
6.6128 | -102.8608
6.6164 | -102.9252
6.62 | -102.9889
6.6236 | -103.0483
6.6272 | -103.1077
6.6308 | -103.1707
6.6344 | -103.2311
6.6381 | -103.2911
6.6417 | -103.3409
6.6453 | -103.3756
6.6489 | -103.402
6.6525 | -103.4236
6.6561 | -103.4321
6.6597 | -103.4244
6.6634 | -103.4081
6.667 | -103.3782
6.6706 | -103.3342
6.6742 | -103.2839
6.6778 | -103.2263
6.6814 | -103.1683
6.685 | -103.1153
6.6886 | -103.066
6.6923 | -103.024
6.6959 | -102.9947
6.6995 | -102.9698
6.7031 | -102.9486
6.7067 | -102.9345
6.7103 | -102.9244
6.7139 | -102.9141
6.7176 | -102.903
6.7212 | -102.8942
6.7248 | -102.8808
6.7284 | -102.862
6.732 | -102.8431
6.7356 | -102.8229
6.7392 | -102.8091
6.7429 | -102.7922
6.7465 | -102.7704
6.7501 | -102.7484
6.7537 | -102.7362
6.7573 | -102.7234
6.7609 | -102.705
6.7645 | -102.6894
6.7681 | -102.6786
6.7718 | -102.6659
6.7754 | -102.6574
6.779 | -102.6559
6.7826 | -102.6536
6.7862 | -102.6474
6.7898 | -102.644
6.7934 | -102.6438
6.7971 | -102.6454
6.8007 | -102.6481
6.8043 | -102.6527
6.8079 | -102.6545
6.8115 | -102.6547
6.8151 | -102.6578
6.8187 | -102.6608
6.8224 | -102.6645
6.826 | -102.6688
6.8296 | -102.6706
6.8332 | -102.6728
6.8368 | -102.6698
6.8404 | -102.6606
6.844 | -102.6531
6.8476 | -102.6433
6.8513 | -102.6322
6.8549 | -102.6194
6.8585 | -102.6062
6.8621 | -102.5958
6.8657 | -102.5855
6.8693 | -102.5749
6.8729 | -102.5743
6.8766 | -102.5753
6.8802 | -102.5738
6.8838 | -102.576
6.8874 | -102.5884
6.891 | -102.6036
6.8946 | -102.6207
6.8982 | -102.6347
6.9018 | -102.6535
6.9055 | -102.6792
6.9091 | -102.7043
6.9127 | -102.7378
6.9163 | -102.7751
6.9199 | -102.8163
6.9235 | -102.8618
6.9271 | -102.9085
6.9308 | -102.9588
6.9344 | -103.0087
6.938 | -103.0565
6.9416 | -103.1036
6.9452 | -103.1387
6.9488 | -103.1632
6.9524 | -103.1847
6.9561 | -103.1961
6.9597 | -103.1951
6.9633 | -103.192
6.9669 | -103.1854
6.9705 | -103.1735
6.9741 | -103.164
6.9777 | -103.1538
6.9813 | -103.1511
6.985 | -103.1547
6.9886 | -103.165
6.9922 | -103.1836
6.9958 | -103.213
6.9994 | -103.2534
7.003 | -103.3007
7.0066 | -103.3497
7.0103 | -103.401
7.0139 | -103.4542
7.0175 | -103.4983
7.0211 | -103.5371
7.0247 | -103.564
7.0283 | -103.577
7.0319 | -103.582
7.0355 | -103.581
7.0392 | -103.5709
7.0428 | -103.5457
7.0464 | -103.5143
7.05 | -103.4824
7.0536 | -103.4422
7.0572 | -103.3957
7.0608 | -103.351
7.0645 | -103.3135
7.0681 | -103.2724
7.0717 | -103.2376
7.0753 | -103.2034
7.0789 | -103.1711
7.0825 | -103.1476
7.0861 | -103.1304
7.0898 | -103.119
7.0934 | -103.113
7.097 | -103.1103
7.1006 | -103.1111
7.1042 | -103.1142
7.1078 | -103.1253
7.1114 | -103.1387
7.115 | -103.1552
7.1187 | -103.1733
7.1223 | -103.1879
7.1259 | -103.1945
7.1295 | -103.2037
7.1331 | -103.2114
7.1367 | -103.2118
7.1403 | -103.2058
7.144 | -103.1955
7.1476 | -103.1846
7.1512 | -103.1683
7.1548 | -103.1539
7.1584 | -103.1473
7.162 | -103.1403
7.1656 | -103.1295
7.1692 | -103.1227
7.1729 | -103.1264
7.1765 | -103.1302
7.1801 | -103.1402
7.1837 | -103.1534
7.1873 | -103.166
7.1909 | -103.1849
7.1945 | -103.2069
7.1982 | -103.2286
7.2018 | -103.2505
7.2054 | -103.2789
7.209 | -103.3064
7.2126 | -103.3311
7.2162 | -103.3571
7.2198 | -103.3806
7.2235 | -103.4039
7.2271 | -103.4252
7.2307 | -103.4515
7.2343 | -103.472
7.2379 | -103.4938
7.2415 | -103.5108
7.2451 | -103.5242
7.2487 | -103.5353
7.2524 | -103.5362
7.256 | -103.5333
7.2596 | -103.5255
7.2632 | -103.5077
7.2668 | -103.4862
7.2704 | -103.4673
7.274 | -103.4444
7.2777 | -103.4236
7.2813 | -103.4084
7.2849 | -103.3982
7.2885 | -103.3953
7.2921 | -103.4022
7.2957 | -103.4347
7.2993 | -103.4506
7.303 | -103.4736
7.3066 | -103.5069
7.3102 | -103.5461
7.3138 | -103.585
7.3174 | -103.623
7.321 | -103.6605
7.3246 | -103.7008
7.3282 | -103.7396
7.3319 | -103.7672
7.3355 | -103.7871
7.3391 | -103.7951
7.3427 | -103.795
7.3463 | -103.7861
7.3499 | -103.7709
7.3535 | -103.7537
7.3572 | -103.7327
7.3608 | -103.7142
7.3644 | -103.6927
7.368 | -103.6687
7.3716 | -103.6424
7.3752 | -103.6223
7.3788 | -103.6043
7.3824 | -103.5846
7.3861 | -103.5675
7.3897 | -103.5547
7.3933 | -103.5439
7.3969 | -103.5336
7.4005 | -103.5261
7.4041 | -103.5232
7.4077 | -103.5243
7.4114 | -103.5213
7.415 | -103.5177
7.4186 | -103.5176
7.4222 | -103.5211
7.4258 | -103.5236
7.4294 | -103.53
7.433 | -103.5325
7.4367 | -103.5291
7.4403 | -103.5225
7.4439 | -103.5179
7.4475 | -103.5136
7.4511 | -103.5059
7.4547 | -103.4936
7.4583 | -103.4751
7.4619 | -103.4608
7.4656 | -103.447
7.4692 | -103.4367
7.4728 | -103.4299
7.4764 | -103.4311
7.48 | -103.4411
7.4836 | -103.4641
7.4872 | -103.4934
7.4909 | -103.5281
7.4945 | -103.5737
7.4981 | -103.623
7.5017 | -103.6761
7.5053 | -103.7339
7.5089 | -103.7856
7.5125 | -103.8332
7.5161 | -103.8804
7.5198 | -103.92
7.5234 | -103.9477
7.527 | -103.9671
7.5306 | -103.9621
7.5342 | -103.9412
7.5378 | -103.915
7.5414 | -103.8816
7.5451 | -103.8409
7.5487 | -103.798
7.5523 | -103.7585
7.5559 | -103.719
7.5595 | -103.6812
7.5631 | -103.6483
7.5667 | -103.6256
7.5704 | -103.6173
7.574 | -103.6135
7.5776 | -103.6147
7.5812 | -103.6283
7.5848 | -103.6475
7.5884 | -103.6712
7.592 | -103.6942
7.5956 | -103.7185
7.5993 | -103.7431
7.6029 | -103.7715
7.6065 | -103.7975
7.6101 | -103.8199
7.6137 | -103.839
7.6173 | -103.85
7.6209 | -103.8646
7.6246 | -103.8696
7.6282 | -103.8683
7.6318 | -103.862
7.6354 | -103.8597
7.639 | -103.8583
7.6426 | -103.8614
7.6462 | -103.8622
7.6498 | -103.8568
7.6535 | -103.853
7.6571 | -103.8533
7.6607 | -103.8485
7.6643 | -103.8408
7.6679 | -103.8262
7.6715 | -103.8103
7.6751 | -103.7992
7.6788 | -103.776
7.6824 | -103.7513
7.686 | -103.722
7.6896 | -103.6858
7.6932 | -103.6443
7.6968 | -103.6042
7.7004 | -103.5659
7.7041 | -103.5225
7.7077 | -103.4754
7.7113 | -103.435
7.7149 | -103.3972
7.7185 | -103.3636
7.7221 | -103.3379
7.7257 | -103.3205
7.7293 | -103.3151
7.733 | -103.3155
7.7366 | -103.3216
7.7402 | -103.3335
7.7438 | -103.3543
7.7474 | -103.3795
7.751 | -103.4043
7.7546 | -103.4335
7.7583 | -103.4709
7.7619 | -103.5171
7.7655 | -103.5628
7.7691 | -103.6105
7.7727 | -103.6617
7.7763 | -103.7098
7.7799 | -103.762
7.7835 | -103.8165
7.7872 | -103.8605
7.7908 | -103.8992
7.7944 | -103.9243
7.798 | -103.9458
7.8016 | -103.9636
7.8052 | -103.967
7.8088 | -103.9573
7.8125 | -103.9384
7.8161 | -103.9236
7.8197 | -103.8942
7.8233 | -103.8812
7.8269 | -103.8608
7.8305 | -103.8399
7.8341 | -103.8238
7.8378 | -103.8136
7.8414 | -103.8118
7.845 | -103.8122
7.8486 | -103.8215
7.8522 | -103.8399
7.8558 | -103.8641
7.8594 | -103.8926
7.863 | -103.9278
7.8667 | -103.9696
7.8703 | -104.0134
7.8739 | -104.0565
7.8775 | -104.0922
7.8811 | -104.1237
7.8847 | -104.1496
7.8883 | -104.1699
7.892 | -104.1869
7.8956 | -104.1979
7.8992 | -104.1922
7.9028 | -104.1852
7.9064 | -104.1779
7.91 | -104.1666
7.9136 | -104.1485
7.9173 | -104.1249
7.9209 | -104.1018
7.9245 | -104.0815
7.9281 | -104.0608
7.9317 | -104.0387
7.9353 | -104.0175
7.9389 | -103.9958
7.9425 | -103.9747
7.9462 | -103.954
7.9498 | -103.9306
7.9534 | -103.9015
7.957 | -103.8718
7.9606 | -103.8438
7.9642 | -103.821
7.9678 | -103.7952
7.9715 | -103.7722
7.9751 | -103.7513
7.9787 | -103.7363
7.9823 | -103.7221
7.9859 | -103.7141
7.9895 | -103.7115
7.9931 | -103.7154
7.9967 | -103.7292
8.0004 | -103.7465
8.004 | -103.7717
8.0076 | -103.7983
8.0112 | -103.8261
8.0148 | -103.8531
8.0184 | -103.8819
8.022 | -103.9087
8.0257 | -103.9349
8.0293 | -103.957
8.0329 | -103.9707
8.0365 | -103.9838
8.0401 | -103.998
8.0437 | -104.0066
8.0473 | -104.009
8.051 | -104.0115
8.0546 | -104.0177
8.0582 | -104.0286
8.0618 | -104.0415
8.0654 | -104.057
8.069 | -104.0738
8.0726 | -104.0958
8.0762 | -104.1169
8.0799 | -104.1299
8.0835 | -104.1411
8.0871 | -104.152
8.0907 | -104.1624
8.0943 | -104.1687
8.0979 | -104.171
8.1015 | -104.1707
8.1052 | -104.1673
8.1088 | -104.1611
8.1124 | -104.1538
8.116 | -104.1407
8.1196 | -104.1266
8.1232 | -104.1071
8.1268 | -104.0927
8.1304 | -104.0763
8.1341 | -104.0581
8.1377 | -104.039
8.1413 | -104.0157
8.1449 | -103.9896
8.1485 | -103.9625
8.1521 | -103.9399
8.1557 | -103.9155
8.1594 | -103.8905
8.163 | -103.8677
8.1666 | -103.8554
8.1702 | -103.8447
8.1738 | -103.8378
8.1774 | -103.8374
8.181 | -103.8463
8.1847 | -103.8605
8.1883 | -103.8837
8.1919 | -103.9091
8.1955 | -103.9388
8.1991 | -103.9748
8.2027 | -104.0077
8.2063 | -104.0444
8.2099 | -104.0804
8.2136 | -104.1157
8.2172 | -104.1483
8.2208 | -104.1798
8.2244 | -104.2057
8.228 | -104.2303
8.2316 | -104.2476
8.2352 | -104.2665
8.2389 | -104.2789
8.2425 | -104.292
8.2461 | -104.2975
8.2497 | -104.2989
8.2533 | -104.2993
8.2569 | -104.296
8.2605 | -104.286
8.2641 | -104.272
8.2678 | -104.2542
8.2714 | -104.2307
8.275 | -104.2049
8.2786 | -104.181
8.2822 | -104.1582
8.2858 | -104.1372
8.2894 | -104.1257
8.2931 | -104.1155
8.2967 | -104.1088
8.3003 | -104.104
8.3039 | -104.1012
8.3075 | -104.1013
8.3111 | -104.1035
8.3147 | -104.1068
8.3184 | -104.1087
8.322 | -104.1149
8.3256 | -104.1181
8.3292 | -104.1184
8.3328 | -104.1217
8.3364 | -104.1243
8.34 | -104.1345
8.3436 | -104.1437
8.3473 | -104.1601
8.3509 | -104.1789
8.3545 | -104.2025
8.3581 | -104.2264
8.3617 | -104.251
8.3653 | -104.2783
8.3689 | -104.3039
8.3726 | -104.3271
8.3762 | -104.348
8.3798 | -104.3667
8.3834 | -104.3857
8.387 | -104.4072
8.3906 | -104.4205
8.3942 | -104.4286
8.3979 | -104.4313
8.4015 | -104.4255
8.4051 | -104.4142
8.4087 | -104.3954
8.4123 | -104.3657
8.4159 | -104.3316
8.4195 | -104.2919
8.4231 | -104.2528
8.4268 | -104.2165
8.4304 | -104.1773
8.434 | -104.1393
8.4376 | -104.105
8.4412 | -104.0778
8.4448 | -104.0534
8.4484 | -104.0371
8.4521 | -104.0288
8.4557 | -104.0242
8.4593 | -104.0302
8.4629 | -104.043
8.4665 | -104.0579
8.4701 | -104.0758
8.4737 | -104.0923
8.4773 | -104.1074
8.481 | -104.1202
8.4846 | -104.1368
8.4882 | -104.1514
8.4918 | -104.1557
8.4954 | -104.1616
8.499 | -104.1649
8.5026 | -104.1597
8.5063 | -104.1506
8.5099 | -104.1391
8.5135 | -104.1227
8.5171 | -104.1061
8.5207 | -104.0879
8.5243 | -104.0735
8.5279 | -104.057
8.5316 | -104.0403
8.5352 | -104.0297
8.5388 | -104.0193
8.5424 | -104.0106
8.546 | -104.0048
8.5496 | -103.9972
8.5532 | -103.9911
8.5568 | -103.988
8.5605 | -103.9888
8.5641 | -103.9882
8.5677 | -103.9877
8.5713 | -103.9895
8.5749 | -103.992
8.5785 | -103.9947
8.5821 | -103.9981
8.5858 | -104.0003
8.5894 | -104.0013
8.593 | -104.0026
8.5966 | -104.0027
8.6002 | -104.0006
8.6038 | -103.9973
8.6074 | -103.9936
8.611 | -103.9778
8.6147 | -103.9583
8.6183 | -103.9402
8.6219 | -103.9217
8.6255 | -103.9039
8.6291 | -103.8889
8.6327 | -103.8758
8.6363 | -103.86
8.64 | -103.8436
8.6436 | -103.8344
8.6472 | -103.8347
8.6508 | -103.8355
8.6544 | -103.8434
8.658 | -103.8607
8.6616 | -103.8786
8.6653 | -103.9093
8.6689 | -103.9514
8.6725 | -103.9997
8.6761 | -104.0544
8.6797 | -104.1163
8.6833 | -104.1852
8.6869 | -104.2592
8.6905 | -104.3364
8.6942 | -104.4098
8.6978 | -104.4802
8.7014 | -104.5423
8.705 | -104.5979
8.7086 | -104.6443
8.7122 | -104.6776
8.7158 | -104.7018
8.7195 | -104.7172
8.7231 | -104.7242
8.7267 | -104.7223
8.7303 | -104.71
8.7339 | -104.6885
8.7375 | -104.6645
8.7411 | -104.638
8.7447 | -104.6038
8.7484 | -104.5657
8.752 | -104.5309
8.7556 | -104.496
8.7592 | -104.465
8.7628 | -104.4435
8.7664 | -104.4249
8.77 | -104.4061
8.7737 | -104.3847
8.7773 | -104.3688
8.7809 | -104.356
8.7845 | -104.3422
8.7881 | -104.3325
8.7917 | -104.3244
8.7953 | -104.3175
8.799 | -104.3116
8.8026 | -104.3082
8.8062 | -104.3089
8.8098 | -104.3088
8.8134 | -104.3164
8.817 | -104.3321
8.8206 | -104.3513
8.8242 | -104.3731
8.8279 | -104.3966
8.8315 | -104.4217
8.8351 | -104.448
8.8387 | -104.4721
8.8423 | -104.4921
8.8459 | -104.5107
8.8495 | -104.5274
8.8532 | -104.5423
8.8568 | -104.5594
8.8604 | -104.567
8.864 | -104.5806
8.8676 | -104.5967
8.8712 | -104.6095
8.8748 | -104.6173
8.8785 | -104.6208
8.8821 | -104.625
8.8857 | -104.6269
8.8893 | -104.6206
8.8929 | -104.6114
8.8965 | -104.6012
8.9001 | -104.5918
8.9037 | -104.5775
8.9074 | -104.5638
8.911 | -104.5496
8.9146 | -104.5351
8.9182 | -104.5186
8.9218 | -104.4974
8.9254 | -104.4742
8.929 | -104.4552
8.9327 | -104.4382
8.9363 | -104.4188
8.9399 | -104.3993
8.9435 | -104.3778
8.9471 | -104.3603
8.9507 | -104.3476
8.9543 | -104.3343
8.9579 | -104.3174
8.9616 | -104.3015
8.9652 | -104.29
8.9688 | -104.285
8.9724 | -104.284
8.976 | -104.2814
8.9796 | -104.2796
8.9832 | -104.2812
8.9869 | -104.2814
8.9905 | -104.2843
8.9941 | -104.2902
8.9977 | -104.2952
9.0013 | -104.3014
9.0049 | -104.3118
9.0085 | -104.3228
9.0122 | -104.3322
9.0158 | -104.342
9.0194 | -104.3557
9.023 | -104.3725
9.0266 | -104.3879
9.0302 | -104.4049
9.0338 | -104.423
9.0374 | -104.4468
9.0411 | -104.4673
9.0447 | -104.4872
9.0483 | -104.5105
9.0519 | -104.5376
9.0555 | -104.5715
9.0591 | -104.6061
9.0627 | -104.6379
9.0664 | -104.6651
9.07 | -104.6929
9.0736 | -104.7235
9.0772 | -104.7511
9.0808 | -104.775
9.0844 | -104.796
9.088 | -104.8196
9.0916 | -104.8465
9.0953 | -104.8688
9.0989 | -104.89
9.1025 | -104.9043
9.1061 | -104.9172
9.1097 | -104.9368
9.1133 | -104.9546
9.1169 | -104.9688
9.1206 | -104.982
9.1242 | -104.9937
9.1278 | -104.9993
9.1314 | -105.0058
9.135 | -105.0113
9.1386 | -105.0137
9.1422 | -105.0167
9.1459 | -105.0183
9.1495 | -105.0105
9.1531 | -104.9978
9.1567 | -104.982
9.1603 | -104.9613
9.1639 | -104.9388
9.1675 | -104.9117
9.1711 | -104.8801
9.1748 | -104.8465
9.1784 | -104.8088
9.182 | -104.7651
9.1856 | -104.7195
9.1892 | -104.6735
9.1928 | -104.6304
9.1964 | -104.5899
9.2001 | -104.5518
9.2037 | -104.5126
9.2073 | -104.4769
9.2109 | -104.4434
9.2145 | -104.4192
9.2181 | -104.4007
9.2217 | -104.3843
9.2253 | -104.3728
9.229 | -104.3703
9.2326 | -104.3698
9.2362 | -104.3764
9.2398 | -104.389
9.2434 | -104.4032
9.247 | -104.4219
9.2506 | -104.4435
9.2543 | -104.4672
9.2579 | -104.4921
9.2615 | -104.5194
9.2651 | -104.5429
9.2687 | -104.5655
9.2723 | -104.5923
9.2759 | -104.6119
9.2796 | -104.6282
9.2832 | -104.6472
9.2868 | -104.6646
9.2904 | -104.6748
9.294 | -104.6835
9.2976 | -104.6839
9.3012 | -104.6804
9.3048 | -104.6757
9.3085 | -104.6699
9.3121 | -104.6609
9.3157 | -104.6498
9.3193 | -104.631
9.3229 | -104.6104
9.3265 | -104.5898
9.3301 | -104.5672
9.3338 | -104.5473
9.3374 | -104.5324
9.341 | -104.5188
9.3446 | -104.5111
9.3482 | -104.5037
9.3518 | -104.5018
9.3554 | -104.5049
9.359 | -104.5098
9.3627 | -104.5141
9.3663 | -104.5228
9.3699 | -104.5338
9.3735 | -104.546
9.3771 | -104.5541
9.3807 | -104.5598
9.3843 | -104.5711
9.388 | -104.585
9.3916 | -104.595
9.3952 | -104.6008
9.3988 | -104.607
9.4024 | -104.6075
9.406 | -104.6072
9.4096 | -104.6022
9.4133 | -104.5975
9.4169 | -104.5933
9.4205 | -104.5868
9.4241 | -104.5792
9.4277 | -104.574
9.4313 | -104.5657
9.4349 | -104.5594
9.4385 | -104.5519
9.4422 | -104.5369
9.4458 | -104.5272
9.4494 | -104.5252
9.453 | -104.5236
9.4566 | -104.5261
9.4602 | -104.5292
9.4638 | -104.536
9.4675 | -104.5503
9.4711 | -104.5655
9.4747 | -104.5777
9.4783 | -104.591
9.4819 | -104.6077
9.4855 | -104.6274
9.4891 | -104.6476
9.4927 | -104.6679
9.4964 | -104.6936
9.5 | -104.7174
9.5036 | -104.7372
9.5072 | -104.7565
9.5108 | -104.7717
9.5144 | -104.7889
9.518 | -104.8031
9.5217 | -104.8124
9.5253 | -104.8204
9.5289 | -104.829
9.5325 | -104.8398
9.5361 | -104.8518
9.5397 | -104.8628
9.5433 | -104.8732
9.547 | -104.8793
9.5506 | -104.8865
9.5542 | -104.8984
9.5578 | -104.9038
9.5614 | -104.9065
9.565 | -104.9108
9.5686 | -104.9165
9.5722 | -104.9242
9.5759 | -104.9286
9.5795 | -104.9352
9.5831 | -104.9427
9.5867 | -104.935
9.5903 | -104.9247
9.5939 | -104.9151
9.5975 | -104.9081
9.6012 | -104.8951
9.6048 | -104.879
9.6084 | -104.8587
9.612 | -104.8387
9.6156 | -104.8189
9.6192 | -104.792
9.6228 | -104.7669
9.6265 | -104.7432
9.6301 | -104.7204
9.6337 | -104.6979
9.6373 | -104.6734
9.6409 | -104.6566
9.6445 | -104.6418
9.6481 | -104.6255
9.6517 | -104.6118
9.6554 | -104.5991
9.659 | -104.5891
9.6626 | -104.5789
9.6662 | -104.5683
9.6698 | -104.56
9.6734 | -104.5569
9.677 | -104.555
9.6807 | -104.5529
9.6843 | -104.5576
9.6879 | -104.5653
9.6915 | -104.575
9.6951 | -104.5831
9.6987 | -104.5954
9.7023 | -104.6104
9.7059 | -104.6236
9.7096 | -104.6408
9.7132 | -104.6564
9.7168 | -104.6741
9.7204 | -104.6927
9.724 | -104.7152
9.7276 | -104.7412
9.7312 | -104.7633
9.7349 | -104.7833
9.7385 | -104.7985
9.7421 | -104.8182
9.7457 | -104.8338
9.7493 | -104.8461
9.7529 | -104.8584
9.7565 | -104.8669
9.7602 | -104.8803
9.7638 | -104.8869
9.7674 | -104.8892
9.771 | -104.8906
9.7746 | -104.8964
9.7782 | -104.9047
9.7818 | -104.9124
9.7854 | -104.9205
9.7891 | -104.9321
9.7927 | -104.9454
9.7963 | -104.9597
9.7999 | -104.9746
9.8035 | -104.9862
9.8071 | -104.9957
9.8107 | -105.0058
9.8144 | -105.0171
9.818 | -105.0289
9.8216 | -105.0411
9.8252 | -105.0525
9.8288 | -105.0636
9.8324 | -105.0759
9.836 | -105.0919
9.8396 | -105.1089
9.8433 | -105.1188
9.8469 | -105.1339
9.8505 | -105.1472
9.8541 | -105.1601
9.8577 | -105.1713
9.8613 | -105.1839
9.8649 | -105.1952
9.8686 | -105.2065
9.8722 | -105.2097
9.8758 | -105.2169
9.8794 | -105.218
9.883 | -105.2186
9.8866 | -105.218
9.8902 | -105.2148
9.8939 | -105.208
9.8975 | -105.2019
9.9011 | -105.1857
9.9047 | -105.1506
9.9083 | -105.1235
9.9119 | -105.0972
9.9155 | -105.0711
9.9191 | -105.0399
9.9228 | -105.0043
9.9264 | -104.9698
9.93 | -104.9391
9.9336 | -104.9082
9.9372 | -104.8808
9.9408 | -104.8552
9.9444 | -104.8307
9.9481 | -104.8167
9.9517 | -104.8069
9.9553 | -104.7973
9.9589 | -104.7882
9.9625 | -104.7844
9.9661 | -104.7872
9.9697 | -104.7913
9.9733 | -104.7964
9.977 | -104.8052
9.9806 | -104.8185
9.9842 | -104.8328
9.9878 | -104.8533
9.9914 | -104.8716
9.995 | -104.8899
9.9986 | -104.9109
10.0023 | -104.9292
10.0059 | -104.9428
10.0095 | -104.958
10.0131 | -104.9702
10.0167 | -104.9833
10.0203 | -104.9898
10.0239 | -104.9982
10.0276 | -105.0008
10.0312 | -105.0005
10.0348 | -104.9995
10.0384 | -104.9932
10.042 | -104.9856
10.0456 | -104.9807
10.0492 | -104.9754
10.0528 | -104.9678
10.0565 | -104.9607
10.0601 | -104.9492
10.0637 | -104.9377
10.0673 | -104.9232
10.0709 | -104.9101
10.0745 | -104.8997
10.0781 | -104.8885
10.0818 | -104.8788
10.0854 | -104.8688
10.089 | -104.8616
10.0926 | -104.8579
10.0962 | -104.854
10.0998 | -104.8521
10.1034 | -104.8547
10.1071 | -104.8609
10.1107 | -104.8694
10.1143 | -104.8762
10.1179 | -104.8854
10.1215 | -104.8982
10.1251 | -104.9137
10.1287 | -104.9294
10.1323 | -104.9454
10.136 | -104.9557
10.1396 | -104.9653
10.1432 | -104.9729
10.1468 | -104.9823
10.1504 | -104.9907
10.154 | -104.9985
10.1576 | -105.0011
10.1613 | -105.0022
10.1649 | -105.0056
10.1685 | -105.0074
10.1721 | -105.0073
10.1757 | -105.0018
10.1793 | -104.9965
10.1829 | -104.9932
10.1865 | -104.9905
10.1902 | -104.9874
10.1938 | -104.9817
10.1974 | -104.976
10.201 | -104.9736
10.2046 | -104.9711
10.2082 | -104.9666
10.2118 | -104.96
10.2155 | -104.9533
10.2191 | -104.9415
10.2227 | -104.9258
10.2263 | -104.9109
10.2299 | -104.8942
10.2335 | -104.8725
10.2371 | -104.8501
10.2408 | -104.8305
10.2444 | -104.8091
10.248 | -104.7886
10.2516 | -104.773
10.2552 | -104.7635
10.2588 | -104.7535
10.2624 | -104.7485
10.266 | -104.7493
10.2697 | -104.7566
10.2733 | -104.7686
10.2769 | -104.7883
10.2805 | -104.8125
10.2841 | -104.8448
10.2877 | -104.8799
10.2913 | -104.9185
10.295 | -104.9563
10.2986 | -104.9918
10.3022 | -105.0271
10.3058 | -105.0646
10.3094 | -105.0983
10.313 | -105.1234
10.3166 | -105.1452
10.3202 | -105.168
10.3239 | -105.1887
10.3275 | -105.2055
10.3311 | -105.2145
10.3347 | -105.2207
10.3383 | -105.2226
10.3419 | -105.222
10.3455 | -105.217
10.3492 | -105.2073
10.3528 | -105.1974
10.3564 | -105.1915
10.36 | -105.1844
10.3636 | -105.1785
10.3672 | -105.1738
10.3708 | -105.166
10.3745 | -105.1627
10.3781 | -105.1621
10.3817 | -105.1672
10.3853 | -105.1754
10.3889 | -105.1843
10.3925 | -105.1973
10.3961 | -105.2137
10.3997 | -105.2254
10.4034 | -105.2354
10.407 | -105.2487
10.4106 | -105.2603
10.4142 | -105.2679
10.4178 | -105.2668
10.4214 | -105.2654
10.425 | -105.2514
10.4287 | -105.2418
10.4323 | -105.2362
10.4359 | -105.2279
10.4395 | -105.2131
10.4431 | -105.1944
10.4467 | -105.1728
10.4503 | -105.1512
10.454 | -105.1291
10.4576 | -105.1052
10.4612 | -105.0828
10.4648 | -105.0596
10.4684 | -105.0377
10.472 | -105.0156
10.4756 | -104.9969
10.4792 | -104.9799
10.4829 | -104.9653
10.4865 | -104.9551
10.4901 | -104.9497
10.4937 | -104.9459
10.4973 | -104.95
10.5009 | -104.957
10.5045 | -104.9671
10.5082 | -104.9783
10.5118 | -104.9904
10.5154 | -105.0039
10.519 | -105.0205
10.5226 | -105.0418
10.5262 | -105.0606
10.5298 | -105.0789
10.5334 | -105.0971
10.5371 | -105.1119
10.5407 | -105.1247
10.5443 | -105.1379
10.5479 | -105.149
10.5515 | -105.1616
10.5551 | -105.1746
10.5587 | -105.1844
10.5624 | -105.1964
10.566 | -105.2108
10.5696 | -105.2262
10.5732 | -105.239
10.5768 | -105.2481
10.5804 | -105.2601
10.584 | -105.2724
10.5877 | -105.2787
10.5913 | -105.2905
10.5949 | -105.3074
10.5985 | -105.3205
10.6021 | -105.3363
10.6057 | -105.3498
10.6093 | -105.3616
10.6129 | -105.3723
10.6166 | -105.3839
10.6202 | -105.3939
10.6238 | -105.3964
10.6274 | -105.3968
10.631 | -105.3998
10.6346 | -105.4008
10.6382 | -105.4
10.6419 | -105.3972
10.6455 | -105.3905
10.6491 | -105.3867
10.6527 | -105.3781
10.6563 | -105.3671
10.6599 | -105.352
10.6635 | -105.3382
10.6671 | -105.3269
10.6708 | -105.3184
10.6744 | -105.3087
10.678 | -105.2931
10.6816 | -105.2787
10.6852 | -105.265
10.6888 | -105.2495
10.6924 | -105.2365
10.6961 | -105.226
10.6997 | -105.2146
10.7033 | -105.2037
10.7069 | -105.1945
10.7105 | -105.1835
10.7141 | -105.1762
10.7177 | -105.1666
10.7214 | -105.1542
10.725 | -105.1466
10.7286 | -105.1414
10.7322 | -105.1422
10.7358 | -105.1464
10.7394 | -105.153
10.743 | -105.1615
10.7466 | -105.1667
10.7503 | -105.173
10.7539 | -105.1833
10.7575 | -105.194
10.7611 | -105.2094
10.7647 | -105.2252
10.7683 | -105.2416
10.7719 | -105.2578
10.7756 | -105.2765
10.7792 | -105.2994
10.7828 | -105.3215
10.7864 | -105.3376
10.79 | -105.352
10.7936 | -105.3659
10.7972 | -105.3766
10.8008 | -105.3887
10.8045 | -105.3953
10.8081 | -105.3995
10.8117 | -105.4038
10.8153 | -105.4084
10.8189 | -105.4056
10.8225 | -105.3972
10.8261 | -105.3865
10.8298 | -105.3733
10.8334 | -105.3612
10.837 | -105.3464
10.8406 | -105.3327
10.8442 | -105.3207
10.8478 | -105.3058
10.8514 | -105.2887
10.8551 | -105.2734
10.8587 | -105.2569
10.8623 | -105.2379
10.8659 | -105.2259
10.8695 | -105.216
10.8731 | -105.2062
10.8767 | -105.1996
10.8803 | -105.1923
10.884 | -105.1849
10.8876 | -105.1797
10.8912 | -105.1757
10.8948 | -105.1749
10.8984 | -105.1766
10.902 | -105.183
10.9056 | -105.1892
10.9093 | -105.1978
10.9129 | -105.2052
10.9165 | -105.2135
10.9201 | -105.223
10.9237 | -105.2326
10.9273 | -105.2446
10.9309 | -105.2612
10.9345 | -105.2802
10.9382 | -105.2995
10.9418 | -105.321
10.9454 | -105.3442
10.949 | -105.3648
10.9526 | -105.385
10.9562 | -105.4029
10.9598 | -105.42
10.9635 | -105.4372
10.9671 | -105.4554
10.9707 | -105.4675
10.9743 | -105.4787
10.9779 | -105.4951
10.9815 | -105.5067
10.9851 | -105.5129
10.9888 | -105.5207
10.9924 | -105.5245
10.996 | -105.5246
10.9996 | -105.5226
11.0032 | -105.5197
11.0068 | -105.5148
11.0104 | -105.5063
11.014 | -105.5
11.0177 | -105.4949
11.0213 | -105.4873
11.0249 | -105.4715
11.0285 | -105.4547
11.0321 | -105.4385
11.0357 | -105.4245
11.0393 | -105.4144
11.043 | -105.4036
11.0466 | -105.3889
11.0502 | -105.3781
11.0538 | -105.3684
11.0574 | -105.3592
11.061 | -105.3479
11.0646 | -105.3379
11.0683 | -105.3293
11.0719 | -105.3228
11.0755 | -105.3157
11.0791 | -105.3099
11.0827 | -105.3068
11.0863 | -105.304
11.0899 | -105.2999
11.0935 | -105.2974
11.0972 | -105.2997
11.1008 | -105.3061
11.1044 | -105.3152
11.108 | -105.3283
11.1116 | -105.3424
11.1152 | -105.3542
11.1188 | -105.3712
11.1225 | -105.3868
11.1261 | -105.4054
11.1297 | -105.426
11.1333 | -105.4501
11.1369 | -105.4719
11.1405 | -105.4895
11.1441 | -105.5095
11.1477 | -105.5253
11.1514 | -105.5346
11.155 | -105.5412
11.1586 | -105.5433
11.1622 | -105.5384
11.1658 | -105.5299
11.1694 | -105.519
11.173 | -105.5056
11.1767 | -105.4824
11.1803 | -105.459
11.1839 | -105.4375
11.1875 | -105.4166
11.1911 | -105.3893
11.1947 | -105.3592
11.1983 | -105.3256
11.202 | -105.2923
11.2056 | -105.2598
11.2092 | -105.2276
11.2128 | -105.1911
11.2164 | -105.1578
11.22 | -105.1282
11.2236 | -105.0999
11.2272 | -105.0754
11.2309 | -105.0515
11.2345 | -105.0236
11.2381 | -104.9976
11.2417 | -104.9743
11.2453 | -104.9558
11.2489 | -104.9384
11.2525 | -104.9256
11.2562 | -104.9111
11.2598 | -104.9006
11.2634 | -104.8957
11.267 | -104.8939
11.2706 | -104.8936
11.2742 | -104.8987
11.2778 | -104.9082
11.2814 | -104.917
11.2851 | -104.9306
11.2887 | -104.9442
11.2923 | -104.9578
11.2959 | -104.9742
11.2995 | -104.9928
11.3031 | -105.0147
11.3067 | -105.0369
11.3104 | -105.0625
11.314 | -105.0887
11.3176 | -105.1132
11.3212 | -105.1375
11.3248 | -105.1629
11.3284 | -105.1847
11.332 | -105.207
11.3357 | -105.2257
11.3393 | -105.2452
11.3429 | -105.2597
11.3465 | -105.268
11.3501 | -105.2771
11.3537 | -105.2877
11.3573 | -105.2954
11.3609 | -105.2967
11.3646 | -105.2897
11.3682 | -105.2852
11.3718 | -105.2836
11.3754 | -105.2773
11.379 | -105.2663
11.3826 | -105.2534
11.3862 | -105.2429
11.3899 | -105.2298
11.3935 | -105.2159
11.3971 | -105.2103
11.4007 | -105.209
11.4043 | -105.2039
11.4079 | -105.1974
11.4115 | -105.1949
11.4152 | -105.1907
11.4188 | -105.1929
11.4224 | -105.1952
11.426 | -105.1983
11.4296 | -105.2031
11.4332 | -105.2126
11.4368 | -105.2228
11.4404 | -105.233
11.4441 | -105.2446
11.4477 | -105.2588
11.4513 | -105.2761
11.4549 | -105.2948
11.4585 | -105.3133
11.4621 | -105.3327
11.4657 | -105.3564
11.4694 | -105.3768
11.473 | -105.3979
11.4766 | -105.4196
11.4802 | -105.441
11.4838 | -105.4646
11.4874 | -105.4905
11.491 | -105.5122
11.4946 | -105.5334
11.4983 | -105.5544
11.5019 | -105.5776
11.5055 | -105.5975
11.5091 | -105.6154
11.5127 | -105.6352
11.5163 | -105.6504
11.5199 | -105.6605
11.5236 | -105.6725
11.5272 | -105.6844
11.5308 | -105.6939
11.5344 | -105.6963
11.538 | -105.6966
11.5416 | -105.6948
11.5452 | -105.6905
11.5488 | -105.6837
11.5525 | -105.6703
11.5561 | -105.6618
11.5597 | -105.6536
11.5633 | -105.6416
11.5669 | -105.6289
11.5705 | -105.6198
11.5741 | -105.6098
11.5778 | -105.5971
11.5814 | -105.5878
11.585 | -105.5762
11.5886 | -105.5599
11.5922 | -105.5437
11.5958 | -105.5275
11.5994 | -105.5127
11.6031 | -105.4987
11.6067 | -105.4861
11.6103 | -105.4751
11.6139 | -105.461
11.6175 | -105.4477
11.6211 | -105.4356
11.6247 | -105.4231
11.6283 | -105.4146
11.632 | -105.4096
11.6356 | -105.4047
11.6392 | -105.4036
11.6428 | -105.4064
11.6464 | -105.4083
11.65 | -105.4114
11.6536 | -105.416
11.6573 | -105.4205
11.6609 | -105.429
11.6645 | -105.4394
11.6681 | -105.4513
11.6717 | -105.4645
11.6753 | -105.4818
11.6789 | -105.5026
11.6826 | -105.5229
11.6862 | -105.5438
11.6898 | -105.5669
11.6934 | -105.5894
11.697 | -105.6158
11.7006 | -105.6441
11.7042 | -105.6709
11.7078 | -105.6958
11.7115 | -105.7219
11.7151 | -105.7434
11.7187 | -105.7616
11.7223 | -105.7786
11.7259 | -105.7893
11.7295 | -105.7962
11.7331 | -105.7994
11.7368 | -105.7959
11.7404 | -105.7839
11.744 | -105.7688
11.7476 | -105.7509
11.7512 | -105.7264
11.7548 | -105.6999
11.7584 | -105.6697
11.762 | -105.6344
11.7657 | -105.6001
11.7693 | -105.5624
11.7729 | -105.5218
11.7765 | -105.4812
11.7801 | -105.4445
11.7837 | -105.4109
11.7873 | -105.3776
11.791 | -105.3466
11.7946 | -105.319
11.7982 | -105.2926
11.8018 | -105.2661
11.8054 | -105.2459
11.809 | -105.2294
11.8126 | -105.2179
11.8163 | -105.211
11.8199 | -105.2072
11.8235 | -105.2056
11.8271 | -105.2081
11.8307 | -105.2131
11.8343 | -105.2192
11.8379 | -105.2273
11.8415 | -105.2364
11.8452 | -105.2495
11.8488 | -105.2649
11.8524 | -105.2804
11.856 | -105.2951
11.8596 | -105.308
11.8632 | -105.3194
11.8668 | -105.329
11.8705 | -105.3387
11.8741 | -105.3496
11.8777 | -105.3605
11.8813 | -105.3706
11.8849 | -105.3785
11.8885 | -105.3838
11.8921 | -105.3868
11.8957 | -105.3848
11.8994 | -105.3816
11.903 | -105.38
11.9066 | -105.3783
11.9102 | -105.3786
11.9138 | -105.3776
11.9174 | -105.3754
11.921 | -105.3685
11.9247 | -105.3627
11.9283 | -105.3588
11.9319 | -105.3541
11.9355 | -105.3446
11.9391 | -105.3323
11.9427 | -105.3222
11.9463 | -105.3126
11.95 | -105.2992
11.9536 | -105.2871
11.9572 | -105.2772
11.9608 | -105.2684
11.9644 | -105.2592
11.968 | -105.2507
11.9716 | -105.2451
11.9752 | -105.235
11.9789 | -105.2255
11.9825 | -105.2159
11.9861 | -105.208
11.9897 | -105.2076
11.9933 | -105.2047
11.9969 | -105.2036
12.0005 | -105.2029
12.0042 | -105.2069
12.0078 | -105.2123
12.0114 | -105.2178
12.015 | -105.2233
12.0186 | -105.2276
12.0222 | -105.2326
12.0258 | -105.2406
12.0295 | -105.2533
12.0331 | -105.268
12.0367 | -105.2786
12.0403 | -105.2888
12.0439 | -105.2995
12.0475 | -105.3122
12.0511 | -105.323
12.0547 | -105.3338
12.0584 | -105.3472
12.062 | -105.3606
12.0656 | -105.3727
12.0692 | -105.3829
12.0728 | -105.3947
12.0764 | -105.4032
12.08 | -105.4127
12.0837 | -105.4226
12.0873 | -105.4315
12.0909 | -105.4401
12.0945 | -105.4503
12.0981 | -105.4569
12.1017 | -105.4637
12.1053 | -105.468
12.1089 | -105.4746
12.1126 | -105.4827
12.1162 | -105.4899
12.1198 | -105.4955
12.1234 | -105.502
12.127 | -105.5109
12.1306 | -105.5212
12.1342 | -105.5322
12.1379 | -105.5439
12.1415 | -105.555
12.1451 | -105.5648
12.1487 | -105.5763
12.1523 | -105.5872
12.1559 | -105.5967
12.1595 | -105.608
12.1632 | -105.6197
12.1668 | -105.6312
12.1704 | -105.6411
12.174 | -105.6489
12.1776 | -105.654
12.1812 | -105.6548
12.1848 | -105.6556
12.1884 | -105.6579
12.1921 | -105.6625
12.1957 | -105.6624
12.1993 | -105.6613
12.2029 | -105.6601
12.2065 | -105.6575
12.2101 | -105.6519
12.2137 | -105.6475
12.2174 | -105.6391
12.221 | -105.6292
12.2246 | -105.6176
12.2282 | -105.6072
12.2318 | -105.5968
12.2354 | -105.5833
12.239 | -105.5677
12.2426 | -105.5475
12.2463 | -105.5313
12.2499 | -105.5118
12.2535 | -105.4955
12.2571 | -105.4808
12.2607 | -105.4659
12.2643 | -105.4533
12.2679 | -105.4403
12.2716 | -105.4297
12.2752 | -105.4209
12.2788 | -105.4113
12.2824 | -105.403
12.286 | -105.395
12.2896 | -105.3908
12.2932 | -105.3888
12.2969 | -105.3868
12.3005 | -105.3841
12.3041 | -105.3803
12.3077 | -105.3783
12.3113 | -105.377
12.3149 | -105.3788
12.3185 | -105.3827
12.3221 | -105.3876
12.3258 | -105.3958
12.3294 | -105.4048
12.333 | -105.4148
12.3366 | -105.4243
12.3402 | -105.434
12.3438 | -105.4445
12.3474 | -105.4592
12.3511 | -105.4755
12.3547 | -105.4926
12.3583 | -105.513
12.3619 | -105.5329
12.3655 | -105.5511
12.3691 | -105.5695
12.3727 | -105.5887
12.3763 | -105.6071
12.38 | -105.6212
12.3836 | -105.6315
12.3872 | -105.6427
12.3908 | -105.6518
12.3944 | -105.6611
12.398 | -105.6682
12.4016 | -105.6698
12.4053 | -105.6677
12.4089 | -105.6642
12.4125 | -105.6603
12.4161 | -105.6534
12.4197 | -105.648
12.4233 | -105.6341
12.4269 | -105.6154
12.4306 | -105.5947
12.4342 | -105.573
12.4378 | -105.5514
12.4414 | -105.529
12.445 | -105.5051
12.4486 | -105.4814
12.4522 | -105.4601
12.4558 | -105.4438
12.4595 | -105.4265
12.4631 | -105.4081
12.4667 | -105.3892
12.4703 | -105.3728
12.4739 | -105.3621
12.4775 | -105.3561
12.4811 | -105.3515
12.4848 | -105.3485
12.4884 | -105.3505
12.492 | -105.3558
12.4956 | -105.3627
12.4992 | -105.3744
12.5028 | -105.388
12.5064 | -105.4033
12.51 | -105.4361
12.5137 | -105.4612
12.5173 | -105.4856
12.5209 | -105.5119
12.5245 | -105.5412
12.5281 | -105.5674
12.5317 | -105.5943
12.5353 | -105.6216
12.539 | -105.648
12.5426 | -105.6804
12.5462 | -105.7096
12.5498 | -105.7378
12.5534 | -105.7637
12.557 | -105.7847
12.5606 | -105.8058
12.5643 | -105.8259
12.5679 | -105.8414
12.5715 | -105.8522
12.5751 | -105.8555
12.5787 | -105.8567
12.5823 | -105.8582
12.5859 | -105.8583
12.5895 | -105.8551
12.5932 | -105.8479
12.5968 | -105.8394
12.6004 | -105.8288
12.604 | -105.8189
12.6076 | -105.8096
12.6112 | -105.7973
12.6148 | -105.7817
12.6185 | -105.7673
12.6221 | -105.7528
12.6257 | -105.7403
12.6293 | -105.7302
12.6329 | -105.7201
12.6365 | -105.7082
12.6401 | -105.6988
12.6438 | -105.6957
12.6474 | -105.6965
12.651 | -105.697
12.6546 | -105.6967
12.6582 | -105.6989
12.6618 | -105.7063
12.6654 | -105.7166
12.669 | -105.7274
12.6727 | -105.7376
12.6763 | -105.7515
12.6799 | -105.7667
12.6835 | -105.7847
12.6871 | -105.8061
12.6907 | -105.8251
12.6943 | -105.8421
12.698 | -105.8632
12.7016 | -105.8868
12.7052 | -105.9101
12.7088 | -105.9349
12.7124 | -105.9594
12.716 | -105.9797
12.7196 | -106.0016
12.7232 | -106.021
12.7269 | -106.0362
12.7305 | -106.0536
12.7341 | -106.0669
12.7377 | -106.076
12.7413 | -106.0849
12.7449 | -106.0909
12.7485 | -106.0973
12.7522 | -106.0994
12.7558 | -106.0944
12.7594 | -106.0902
12.763 | -106.0786
12.7666 | -106.0634
12.7702 | -106.0429
12.7738 | -106.0202
12.7775 | -105.9927
12.7811 | -105.963
12.7847 | -105.9369
12.7883 | -105.9086
12.7919 | -105.8818
12.7955 | -105.8544
12.7991 | -105.8277
12.8027 | -105.8
12.8064 | -105.776
12.81 | -105.7553
12.8136 | -105.7334
12.8172 | -105.7123
12.8208 | -105.6932
12.8244 | -105.68
12.828 | -105.6679
12.8317 | -105.6585
12.8353 | -105.653
12.8389 | -105.647
12.8425 | -105.6422
12.8461 | -105.6416
12.8497 | -105.646
12.8533 | -105.6514
12.8569 | -105.6602
12.8606 | -105.6708
12.8642 | -105.6835
12.8678 | -105.6974
12.8714 | -105.7114
12.875 | -105.7293
12.8786 | -105.7494
12.8822 | -105.7687
12.8859 | -105.787
12.8895 | -105.8032
12.8931 | -105.8211
12.8967 | -105.8438
12.9003 | -105.8692
12.9039 | -105.8927
12.9075 | -105.9131
12.9112 | -105.9313
12.9148 | -105.9495
12.9184 | -105.9701
12.922 | -105.9955
12.9256 | -106.0169
12.9292 | -106.034
12.9328 | -106.047
12.9364 | -106.062
12.9401 | -106.0802
12.9437 | -106.0944
12.9473 | -106.1101
12.9509 | -106.1239
12.9545 | -106.1353
12.9581 | -106.1436
12.9617 | -106.1547
12.9654 | -106.1606
12.969 | -106.1616
12.9726 | -106.1641
12.9762 | -106.1619
12.9798 | -106.1587
12.9834 | -106.1575
12.987 | -106.1557
12.9907 | -106.1493
12.9943 | -106.1396
12.9979 | -106.1299
13.0015 | -106.1206
13.0051 | -106.11
13.0087 | -106.0963
13.0123 | -106.08
13.0159 | -106.0627
13.0196 | -106.0445
13.0232 | -106.0304
13.0268 | -106.0132
13.0304 | -105.9933
13.034 | -105.9757
13.0376 | -105.9565
13.0412 | -105.9396
13.0449 | -105.9247
13.0485 | -105.9092
13.0521 | -105.8957
13.0557 | -105.8811
13.0593 | -105.8704
13.0629 | -105.8604
13.0665 | -105.8493
13.0701 | -105.8415
13.0738 | -105.8385
13.0774 | -105.8377
13.081 | -105.8356
13.0846 | -105.8317
13.0882 | -105.8301
13.0918 | -105.8286
13.0954 | -105.8317
13.0991 | -105.8336
13.1027 | -105.8331
13.1063 | -105.8335
13.1099 | -105.8368
13.1135 | -105.8422
13.1171 | -105.8503
13.1207 | -105.8601
13.1244 | -105.8678
13.128 | -105.8744
13.1316 | -105.8827
13.1352 | -105.8917
13.1388 | -105.9043
13.1424 | -105.9173
13.146 | -105.9306
13.1496 | -105.9427
13.1533 | -105.9566
13.1569 | -105.968
13.1605 | -105.976
13.1641 | -105.984
13.1677 | -105.9943
13.1713 | -106.0008
13.1749 | -106.0079
13.1786 | -106.0154
13.1822 | -106.0238
13.1858 | -106.0307
13.1894 | -106.0355
13.193 | -106.0402
13.1966 | -106.0427
13.2002 | -106.0426
13.2038 | -106.0421
13.2075 | -106.0432
13.2111 | -106.0465
13.2147 | -106.0512
13.2183 | -106.0528
13.2219 | -106.0547
13.2255 | -106.056
13.2291 | -106.0545
13.2328 | -106.0547
13.2364 | -106.0553
13.24 | -106.056
13.2436 | -106.0577
13.2472 | -106.0605
13.2508 | -106.0632
13.2544 | -106.0657
13.2581 | -106.0689
13.2617 | -106.0684
13.2653 | -106.0678
13.2689 | -106.0702
13.2725 | -106.074
13.2761 | -106.0778
13.2797 | -106.0812
13.2833 | -106.0853
13.287 | -106.0888
13.2906 | -106.0932
13.2942 | -106.101
13.2978 | -106.1086
13.3014 | -106.1108
13.305 | -106.1137
13.3086 | -106.1153
13.3123 | -106.1157
13.3159 | -106.1189
13.3195 | -106.1247
13.3231 | -106.1274
13.3267 | -106.1295
13.3303 | -106.1302
13.3339 | -106.1332
13.3375 | -106.1369
13.3412 | -106.139
13.3448 | -106.1399
13.3484 | -106.1408
13.352 | -106.1443
13.3556 | -106.1473
13.3592 | -106.152
13.3628 | -106.1571
13.3665 | -106.1578
13.3701 | -106.1544
13.3737 | -106.1535
13.3773 | -106.1525
13.3809 | -106.1514
13.3845 | -106.1519
13.3881 | -106.1534
13.3918 | -106.1518
13.3954 | -106.1493
13.399 | -106.1466
13.4026 | -106.1379
13.4062 | -106.1281
13.4098 | -106.1209
13.4134 | -106.114
13.417 | -106.1059
13.4207 | -106.0986
13.4243 | -106.0903
13.4279 | -106.0796
13.4315 | -106.0715
13.4351 | -106.0643
13.4387 | -106.0583
13.4423 | -106.0488
13.446 | -106.042
13.4496 | -106.0362
13.4532 | -106.0324
13.4568 | -106.0288
13.4604 | -106.0238
13.464 | -106.0188
13.4676 | -106.0139
13.4712 | -106.0112
13.4749 | -106.0122
13.4785 | -106.0161
13.4821 | -106.0216
13.4857 | -106.028
13.4893 | -106.0344
13.4929 | -106.0436
13.4965 | -106.0526
13.5002 | -106.0642
13.5038 | -106.0755
13.5074 | -106.0851
13.511 | -106.0922
13.5146 | -106.0996
13.5182 | -106.1073
13.5218 | -106.1167
13.5255 | -106.1254
13.5291 | -106.1304
13.5327 | -106.1373
13.5363 | -106.1405
13.5399 | -106.14
13.5435 | -106.1406
13.5471 | -106.1394
13.5507 | -106.1284
13.5544 | -106.1198
13.558 | -106.1112
13.5616 | -106.1061
13.5652 | -106.1017
13.5688 | -106.0947
13.5724 | -106.0865
13.576 | -106.0784
13.5797 | -106.069
13.5833 | -106.0584
13.5869 | -106.0526
13.5905 | -106.0454
13.5941 | -106.0363
13.5977 | -106.032
13.6013 | -106.0282
13.605 | -106.0218
13.6086 | -106.0156
13.6122 | -106.0123
13.6158 | -106.012
13.6194 | -106.009
13.623 | -106.0068
13.6266 | -106.0045
13.6302 | -106.0026
13.6339 | -106.0028
13.6375 | -106.0038
13.6411 | -106.0018
13.6447 | -106.0036
13.6483 | -106.0055
13.6519 | -106.0112
13.6555 | -106.0171
13.6592 | -106.0186
13.6628 | -106.0235
13.6664 | -106.0259
13.67 | -106.0265
13.6736 | -106.0276
13.6772 | -106.0311
13.6808 | -106.0324
13.6844 | -106.0325
13.6881 | -106.0353
13.6917 | -106.0393
13.6953 | -106.0372
13.6989 | -106.0339
13.7025 | -106.0326
13.7061 | -106.0318
13.7097 | -106.0292
13.7134 | -106.0262
13.717 | -106.0201
13.7206 | -106.0144
13.7242 | -106.0141
13.7278 | -106.0116
13.7314 | -106.0111
13.735 | -106.0115
13.7387 | -106.0104
13.7423 | -106.0117
13.7459 | -106.0142
13.7495 | -106.0167
13.7531 | -106.0212
13.7567 | -106.0283
13.7603 | -106.0351
13.7639 | -106.0424
13.7676 | -106.0464
13.7712 | -106.0523
13.7748 | -106.0576
13.7784 | -106.0617
13.782 | -106.0688
13.7856 | -106.0752
13.7892 | -106.08
13.7929 | -106.0826
13.7965 | -106.0827
13.8001 | -106.0828
13.8037 | -106.0841
13.8073 | -106.0819
13.8109 | -106.0748
13.8145 | -106.0687
13.8181 | -106.0646
13.8218 | -106.0595
13.8254 | -106.0523
13.829 | -106.0406
13.8326 | -106.0284
13.8362 | -106.0191
13.8398 | -106.0098
13.8434 | -105.9975
13.8471 | -105.9875
13.8507 | -105.9754
13.8543 | -105.9641
13.8579 | -105.9608
13.8615 | -105.9578
13.8651 | -105.9542
13.8687 | -105.9519
13.8724 | -105.9521
13.876 | -105.9551
13.8796 | -105.9605
13.8832 | -105.9675
13.8868 | -105.9747
13.8904 | -105.9838
13.894 | -105.9992
13.8976 | -106.0152
13.9013 | -106.0277
13.9049 | -106.041
13.9085 | -106.0572
13.9121 | -106.0711
13.9157 | -106.0862
13.9193 | -106.1022
13.9229 | -106.1163
13.9266 | -106.1333
13.9302 | -106.1512
13.9338 | -106.1661
13.9374 | -106.1797
13.941 | -106.192
13.9446 | -106.2041
13.9482 | -106.2161
13.9519 | -106.2206
13.9555 | -106.2221
13.9591 | -106.2192
13.9627 | -106.2119
13.9663 | -106.2046
13.9699 | -106.1971
13.9735 | -106.1906
13.9771 | -106.1821
13.9808 | -106.1665
13.9844 | -106.145
13.988 | -106.122
13.9916 | -106.1018
13.9952 | -106.0797
13.9988 | -106.0545
14.0024 | -106.0283
14.0061 | -106.0008
14.0097 | -105.9732
14.0133 | -105.9419
14.0169 | -105.9108
14.0205 | -105.8825
14.0241 | -105.856
14.0277 | -105.8326
14.0313 | -105.8117
14.035 | -105.7906
14.0386 | -105.7729
14.0422 | -105.7586
14.0458 | -105.746
14.0494 | -105.7347
14.053 | -105.728
14.0566 | -105.7248
14.0603 | -105.7253
14.0639 | -105.7315
14.0675 | -105.7406
14.0711 | -105.7518
14.0747 | -105.7801
14.0783 | -105.7931
14.0819 | -105.8095
14.0855 | -105.8294
14.0892 | -105.8525
14.0928 | -105.8769
14.0964 | -105.8994
14.1 | -105.9251
14.1036 | -105.9534
14.1072 | -105.9791
14.1108 | -106.0002
14.1145 | -106.0231
14.1181 | -106.0463
14.1217 | -106.0665
14.1253 | -106.0842
14.1289 | -106.1001
14.1325 | -106.114
14.1361 | -106.1295
14.1398 | -106.1457
14.1434 | -106.1599
14.147 | -106.1702
14.1506 | -106.1766
14.1542 | -106.1854
14.1578 | -106.1952
14.1614 | -106.2037
14.165 | -106.2122
14.1687 | -106.2165
14.1723 | -106.2166
14.1759 | -106.2144
14.1795 | -106.2138
14.1831 | -106.2167
14.1867 | -106.2208
14.1903 | -106.2256
14.194 | -106.2293
14.1976 | -106.2329
14.2012 | -106.2376
14.2048 | -106.2433
14.2084 | -106.2524
14.212 | -106.2623
14.2156 | -106.274
14.2193 | -106.2846
14.2229 | -106.2943
14.2265 | -106.3043
14.2301 | -106.3164
14.2337 | -106.3262
14.2373 | -106.3356
14.2409 | -106.3457
14.2445 | -106.3578
14.2482 | -106.3687
14.2518 | -106.3794
14.2554 | -106.3862
14.259 | -106.3903
14.2626 | -106.3934
14.2662 | -106.398
14.2698 | -106.4006
14.2735 | -106.398
14.2771 | -106.3919
14.2807 | -106.3879
14.2843 | -106.3826
14.2879 | -106.3744
14.2915 | -106.3633
14.2951 | -106.3512
14.2987 | -106.3371
14.3024 | -106.3239
14.306 | -106.3089
14.3096 | -106.2924
14.3132 | -106.2772
14.3168 | -106.2643
14.3204 | -106.2481
14.324 | -106.23
14.3277 | -106.2116
14.3313 | -106.1913
14.3349 | -106.1722
14.3385 | -106.1572
14.3421 | -106.143
14.3457 | -106.1247
14.3493 | -106.1095
14.353 | -106.1009
14.3566 | -106.0973
14.3602 | -106.0953
14.3638 | -106.093
14.3674 | -106.0931
14.371 | -106.0957
14.3746 | -106.1014
14.3782 | -106.1088
14.3819 | -106.1149
14.3855 | -106.1216
14.3891 | -106.1292
14.3927 | -106.1381
14.3963 | -106.1445
14.3999 | -106.1514
14.4035 | -106.1602
14.4072 | -106.1706
14.4108 | -106.1817
14.4144 | -106.1931
14.418 | -106.2036
14.4216 | -106.2116
14.4252 | -106.2225
14.4288 | -106.2334
14.4324 | -106.2427
14.4361 | -106.2484
14.4397 | -106.2536
14.4433 | -106.258
14.4469 | -106.2597
14.4505 | -106.2591
14.4541 | -106.258
14.4577 | -106.2568
14.4614 | -106.2541
14.465 | -106.2508
14.4686 | -106.2469
14.4722 | -106.2385
14.4758 | -106.2289
14.4794 | -106.2206
14.483 | -106.2094
14.4867 | -106.1936
14.4903 | -106.1784
14.4939 | -106.1636
14.4975 | -106.1448
14.5011 | -106.123
14.5047 | -106.1043
14.5083 | -106.0872
14.5119 | -106.0708
14.5156 | -106.056
14.5192 | -106.0397
14.5228 | -106.0235
14.5264 | -106.0101
14.53 | -105.9992
14.5336 | -105.9942
14.5372 | -105.9913
14.5409 | -105.9851
14.5445 | -105.9783
14.5481 | -105.9759
14.5517 | -105.9767
14.5553 | -105.9781
14.5589 | -105.9788
14.5625 | -105.982
14.5662 | -105.9866
14.5698 | -105.9905
14.5734 | -105.9981
14.577 | -106.0069
14.5806 | -106.0172
14.5842 | -106.0264
14.5878 | -106.0396
14.5914 | -106.051
14.5951 | -106.061
14.5987 | -106.0734
14.6023 | -106.0868
14.6059 | -106.1007
14.6095 | -106.1123
14.6131 | -106.1215
14.6167 | -106.1288
14.6204 | -106.1352
14.624 | -106.1422
14.6276 | -106.1509
14.6312 | -106.1607
14.6348 | -106.1683
14.6384 | -106.1761
14.642 | -106.1874
14.6456 | -106.1993
14.6493 | -106.2112
14.6529 | -106.2232
14.6565 | -106.2354
14.6601 | -106.2473
14.6637 | -106.2599
14.6673 | -106.2739
14.6709 | -106.29
14.6746 | -106.3022
14.6782 | -106.3141
14.6818 | -106.3272
14.6854 | -106.3407
14.689 | -106.356
14.6926 | -106.3738
14.6962 | -106.3898
14.6999 | -106.4045
14.7035 | -106.4212
14.7071 | -106.4426
14.7107 | -106.4609
14.7143 | -106.4754
14.7179 | -106.4888
14.7215 | -106.5007
14.7251 | -106.5115
14.7288 | -106.524
14.7324 | -106.535
14.736 | -106.544
14.7396 | -106.5504
14.7432 | -106.5512
14.7468 | -106.5527
14.7504 | -106.5546
14.7541 | -106.5537
14.7577 | -106.5505
14.7613 | -106.5454
14.7649 | -106.5391
14.7685 | -106.5326
14.7721 | -106.5218
14.7757 | -106.5085
14.7793 | -106.4984
14.783 | -106.4866
14.7866 | -106.4742
14.7902 | -106.46
14.7938 | -106.4484
14.7974 | -106.4352
14.801 | -106.4246
14.8046 | -106.415
14.8083 | -106.405
14.8119 | -106.3938
14.8155 | -106.3824
14.8191 | -106.3736
14.8227 | -106.367
14.8263 | -106.3609
14.8299 | -106.3558
14.8336 | -106.3495
14.8372 | -106.3435
14.8408 | -106.3395
14.8444 | -106.337
14.848 | -106.3332
14.8516 | -106.328
14.8552 | -106.3234
14.8588 | -106.3218
14.8625 | -106.3219
14.8661 | -106.3213
14.8697 | -106.3203
14.8733 | -106.3206
14.8769 | -106.3218
14.8805 | -106.3256
14.8841 | -106.3307
14.8878 | -106.3361
14.8914 | -106.3424
14.895 | -106.349
14.8986 | -106.3543
14.9022 | -106.3591
14.9058 | -106.3655
14.9094 | -106.3763
14.913 | -106.3864
14.9167 | -106.3931
14.9203 | -106.399
14.9239 | -106.4088
14.9275 | -106.4179
14.9311 | -106.428
14.9347 | -106.4404
14.9383 | -106.4516
14.942 | -106.4601
14.9456 | -106.47
14.9492 | -106.4765
14.9528 | -106.4803
14.9564 | -106.4841
14.96 | -106.486
14.9636 | -106.4887
14.9673 | -106.4914
14.9709 | -106.488
14.9745 | -106.4816
14.9781 | -106.4742
14.9817 | -106.4654
14.9853 | -106.4581
14.9889 | -106.4491
14.9925 | -106.4397
14.9962 | -106.426
14.9998 | -106.4091
15.0034 | -106.393
15.007 | -106.3794
15.0106 | -106.3635
15.0142 | -106.3449
15.0178 | -106.3261
15.0215 | -106.3049
15.0251 | -106.2856
15.0287 | -106.2681
15.0323 | -106.2498
15.0359 | -106.2318
15.0395 | -106.2144
15.0431 | -106.1982
15.0467 | -106.1835
15.0504 | -106.1683
15.054 | -106.1514
15.0576 | -106.1329
15.0612 | -106.1191
15.0648 | -106.1085
15.0684 | -106.0974
15.072 | -106.0871
15.0757 | -106.0763
15.0793 | -106.0701
15.0829 | -106.0661
15.0865 | -106.0593
15.0901 | -106.0528
15.0937 | -106.047
15.0973 | -106.0414
15.101 | -106.041
15.1046 | -106.0427
15.1082 | -106.0449
15.1118 | -106.0484
15.1154 | -106.0493
15.119 | -106.0522
15.1226 | -106.0524
15.1262 | -106.0535
15.1299 | -106.0571
15.1335 | -106.0611
15.1371 | -106.0661
15.1407 | -106.0717
15.1443 | -106.0778
15.1479 | -106.0797
15.1515 | -106.0814
15.1552 | -106.0844
15.1588 | -106.0894
15.1624 | -106.093
15.166 | -106.0988
15.1696 | -106.106
15.1732 | -106.1105
15.1768 | -106.1125
15.1805 | -106.1127
15.1841 | -106.1137
15.1877 | -106.1176
15.1913 | -106.122
15.1949 | -106.1245
15.1985 | -106.1266
15.2021 | -106.1279
15.2057 | -106.1311
15.2094 | -106.1299
15.213 | -106.1297
15.2166 | -106.1341
15.2202 | -106.1365
15.2238 | -106.1349
15.2274 | -106.1357
15.231 | -106.1398
15.2347 | -106.1433
15.2383 | -106.1453
15.2419 | -106.1465
15.2455 | -106.1474
15.2491 | -106.1468
15.2527 | -106.1467
15.2563 | -106.1481
15.2599 | -106.1491
15.2636 | -106.1511
15.2672 | -106.1519
15.2708 | -106.1521
15.2744 | -106.1547
15.278 | -106.1581
15.2816 | -106.1592
15.2852 | -106.1602
15.2889 | -106.1606
15.2925 | -106.1614
15.2961 | -106.1615
15.2997 | -106.1637
15.3033 | -106.1688
15.3069 | -106.1756
15.3105 | -106.1816
15.3142 | -106.1871
15.3178 | -106.1897
15.3214 | -106.1935
15.325 | -106.2013
15.3286 | -106.2087
15.3322 | -106.2163
15.3358 | -106.2271
15.3394 | -106.2361
15.3431 | -106.2473
15.3467 | -106.2536
15.3503 | -106.2593
15.3539 | -106.2686
15.3575 | -106.2778
15.3611 | -106.2861
15.3647 | -106.2965
15.3684 | -106.3078
15.372 | -106.3172
15.3756 | -106.327
15.3792 | -106.3368
15.3828 | -106.3441
15.3864 | -106.3503
15.39 | -106.3539
15.3936 | -106.3596
15.3973 | -106.3664
15.4009 | -106.3721
15.4045 | -106.3756
15.4081 | -106.3771
15.4117 | -106.3791
15.4153 | -106.3788
15.4189 | -106.3785
15.4226 | -106.375
15.4262 | -106.37
15.4298 | -106.3687
15.4334 | -106.3648
15.437 | -106.3555
15.4406 | -106.3486
15.4442 | -106.3422
15.4479 | -106.3326
15.4515 | -106.3229
15.4551 | -106.3145
15.4587 | -106.3069
15.4623 | -106.2989
15.4659 | -106.2898
15.4695 | -106.2812
15.4731 | -106.2718
15.4768 | -106.2609
15.4804 | -106.2505
15.484 | -106.2396
15.4876 | -106.2317
15.4912 | -106.2269
15.4948 | -106.2198
15.4984 | -106.2117
15.5021 | -106.2038
15.5057 | -106.1958
15.5093 | -106.1892
15.5129 | -106.1826
15.5165 | -106.1779
15.5201 | -106.1715
15.5237 | -106.166
15.5274 | -106.1594
15.531 | -106.1539
15.5346 | -106.1501
15.5382 | -106.1486
15.5418 | -106.1455
15.5454 | -106.1404
15.549 | -106.1351
15.5526 | -106.1327
15.5563 | -106.128
15.5599 | -106.1237
15.5635 | -106.1196
15.5671 | -106.1204
15.5707 | -106.1193
15.5743 | -106.1147
15.5779 | -106.1106
15.5816 | -106.1084
15.5852 | -106.1049
15.5888 | -106.1004
15.5924 | -106.0965
15.596 | -106.0915
15.5996 | -106.0839
15.6032 | -106.0778
15.6068 | -106.0729
15.6105 | -106.0671
15.6141 | -106.061
15.6177 | -106.0546
15.6213 | -106.0496
15.6249 | -106.0437
15.6285 | -106.0382
15.6321 | -106.0336
15.6358 | -106.031
15.6394 | -106.0306
15.643 | -106.029
15.6466 | -106.0288
15.6502 | -106.0286
15.6538 | -106.027
15.6574 | -106.0291
15.6611 | -106.0316
15.6647 | -106.0302
15.6683 | -106.0289
15.6719 | -106.0293
15.6755 | -106.0293
15.6791 | -106.0299
15.6827 | -106.0325
15.6863 | -106.035
15.69 | -106.0382
15.6936 | -106.0415
15.6972 | -106.0453
15.7008 | -106.0508
15.7044 | -106.0549
15.708 | -106.0569
15.7116 | -106.0596
15.7153 | -106.0646
15.7189 | -106.072
15.7225 | -106.0771
15.7261 | -106.0837
15.7297 | -106.0917
15.7333 | -106.0994
15.7369 | -106.1064
15.7405 | -106.1144
15.7442 | -106.1214
15.7478 | -106.126
15.7514 | -106.1309
15.755 | -106.1381
15.7586 | -106.1447
15.7622 | -106.1488
15.7658 | -106.1531
15.7695 | -106.1571
15.7731 | -106.1635
15.7767 | -106.1672
15.7803 | -106.1684
15.7839 | -106.1691
15.7875 | -106.1694
15.7911 | -106.1661
15.7948 | -106.1652
15.7984 | -106.1655
15.802 | -106.1645
15.8056 | -106.1608
15.8092 | -106.1598
15.8128 | -106.1596
15.8164 | -106.1581
15.82 | -106.1547
15.8237 | -106.1514
15.8273 | -106.1505
15.8309 | -106.1467
15.8345 | -106.1442
15.8381 | -106.1418
15.8417 | -106.1363
15.8453 | -106.1292
15.849 | -106.1214
15.8526 | -106.1153
15.8562 | -106.1098
15.8598 | -106.1055
15.8634 | -106.1022
15.867 | -106.0958
15.8706 | -106.0866
15.8742 | -106.0795
15.8779 | -106.0753
15.8815 | -106.0702
15.8851 | -106.0636
15.8887 | -106.0594
15.8923 | -106.0565
15.8959 | -106.0527
15.8995 | -106.0504
15.9032 | -106.0461
15.9068 | -106.0413
15.9104 | -106.0359
15.914 | -106.0298
15.9176 | -106.025
15.9212 | -106.0192
15.9248 | -106.0106
15.9285 | -106.0032
15.9321 | -105.9975
15.9357 | -105.9947
15.9393 | -105.9898
15.9429 | -105.9833
15.9465 | -105.9776
15.9501 | -105.9747
15.9537 | -105.9726
15.9574 | -105.9682
15.961 | -105.9647
15.9646 | -105.9621
15.9682 | -105.9572
15.9718 | -105.9537
15.9754 | -105.9537
15.979 | -105.9529
15.9827 | -105.9499
15.9863 | -105.9474
15.9899 | -105.9474
15.9935 | -105.9481
15.9971 | -105.9506
16.0007 | -105.9519
16.0043 | -105.9544
16.008 | -105.9577
16.0116 | -105.9592
16.0152 | -105.9627
16.0188 | -105.9668
16.0224 | -105.9693
16.026 | -105.9718
16.0296 | -105.9751
16.0332 | -105.9809
16.0369 | -105.982
16.0405 | -105.985
16.0441 | -105.99
16.0477 | -105.9964
16.0513 | -106.0045
16.0549 | -106.0122
16.0585 | -106.0197
16.0622 | -106.0272
16.0658 | -106.0349
16.0694 | -106.0441
16.073 | -106.0522
16.0766 | -106.0636
16.0802 | -106.075
16.0838 | -106.0851
16.0874 | -106.0934
16.0911 | -106.1013
16.0947 | -106.1127
16.0983 | -106.1265
16.1019 | -106.1398
16.1055 | -106.1498
16.1091 | -106.159
16.1127 | -106.1678
16.1164 | -106.1786
16.12 | -106.1901
16.1236 | -106.2014
16.1272 | -106.2127
16.1308 | -106.225
16.1344 | -106.2357
16.138 | -106.2469
16.1416 | -106.259
16.1453 | -106.2678
16.1489 | -106.2753
16.1525 | -106.2854
16.1561 | -106.294
16.1597 | -106.2958
16.1633 | -106.3006
16.1669 | -106.3038
16.1706 | -106.309
16.1742 | -106.3106
16.1778 | -106.3107
16.1814 | -106.3099
16.185 | -106.3109
16.1886 | -106.3106
16.1922 | -106.3097
16.1959 | -106.3086
16.1995 | -106.3063
16.2031 | -106.3022
16.2067 | -106.2978
16.2103 | -106.2935
16.2139 | -106.2889
16.2175 | -106.2839
16.2211 | -106.2794
16.2248 | -106.2758
16.2284 | -106.2717
16.232 | -106.2688
16.2356 | -106.266
16.2392 | -106.265
16.2428 | -106.2666
16.2464 | -106.2671
16.2501 | -106.2682
16.2537 | -106.2708
16.2573 | -106.2756
16.2609 | -106.2796
16.2645 | -106.283
16.2681 | -106.2931
16.2717 | -106.3037
16.2754 | -106.3136
16.279 | -106.3243
16.2826 | -106.3364
16.2862 | -106.3481
16.2898 | -106.3623
16.2934 | -106.3775
16.297 | -106.3929
16.3006 | -106.4101
16.3043 | -106.4301
16.3079 | -106.4477
16.3115 | -106.4618
16.3151 | -106.4766
16.3187 | -106.4905
16.3223 | -106.5052
16.3259 | -106.5204
16.3296 | -106.5339
16.3332 | -106.5454
16.3368 | -106.5571
16.3404 | -106.5693
16.344 | -106.5819
16.3476 | -106.5883
16.3512 | -106.591
16.3548 | -106.5971
16.3585 | -106.6012
16.3621 | -106.6006
16.3657 | -106.6021
16.3693 | -106.6048
16.3729 | -106.6045
16.3765 | -106.6036
16.3801 | -106.6063
16.3838 | -106.6041
16.3874 | -106.5991
16.391 | -106.5945
16.3946 | -106.5901
16.3982 | -106.5835
16.4018 | -106.5773
16.4054 | -106.5698
16.4091 | -106.5647
16.4127 | -106.5598
16.4163 | -106.5541
16.4199 | -106.5448
16.4235 | -106.5387
16.4271 | -106.5316
16.4307 | -106.5246
16.4343 | -106.5212
16.438 | -106.5205
16.4416 | -106.5168
16.4452 | -106.5134
16.4488 | -106.512
16.4524 | -106.5099
16.456 | -106.5075
16.4596 | -106.5052
16.4633 | -106.5066
16.4669 | -106.5068
16.4705 | -106.5087
16.4741 | -106.5112
16.4777 | -106.5151
16.4813 | -106.5188
16.4849 | -106.5233
16.4886 | -106.5282
16.4922 | -106.5342
16.4958 | -106.5425
16.4994 | -106.5505
16.503 | -106.5539
16.5066 | -106.5555
16.5102 | -106.5595
16.5138 | -106.5644
16.5175 | -106.5662
16.5211 | -106.5709
16.5247 | -106.577
16.5283 | -106.5806
16.5319 | -106.5842
16.5355 | -106.5881
16.5391 | -106.5911
16.5428 | -106.5925
16.5464 | -106.5978
16.55 | -106.6039
16.5536 | -106.6099
16.5572 | -106.6124
16.5608 | -106.6119
16.5644 | -106.6148
16.568 | -106.6176
16.5717 | -106.6191
16.5753 | -106.6207
16.5789 | -106.6213
16.5825 | -106.6234
16.5861 | -106.6256
16.5897 | -106.6209
16.5933 | -106.6162
16.597 | -106.6147
16.6006 | -106.6077
16.6042 | -106.5981
16.6078 | -106.5937
16.6114 | -106.5876
16.615 | -106.5788
16.6186 | -106.5701
16.6223 | -106.5621
16.6259 | -106.5527
16.6295 | -106.5425
16.6331 | -106.5357
16.6367 | -106.5278
16.6403 | -106.5174
16.6439 | -106.5053
16.6475 | -106.4921
16.6512 | -106.4803
16.6548 | -106.4715
16.6584 | -106.4636
16.662 | -106.4566
16.6656 | -106.448
16.6692 | -106.4381
16.6728 | -106.4282
16.6765 | -106.4198
16.6801 | -106.418
16.6837 | -106.4137
16.6873 | -106.4083
16.6909 | -106.4047
16.6945 | -106.4022
16.6981 | -106.3999
16.7017 | -106.3988
16.7054 | -106.4008
16.709 | -106.4015
16.7126 | -106.4043
16.7162 | -106.4078
16.7198 | -106.4111
16.7234 | -106.4165
16.727 | -106.4224
16.7307 | -106.4269
16.7343 | -106.4325
16.7379 | -106.4398
16.7415 | -106.4515
16.7451 | -106.4652
16.7487 | -106.4756
16.7523 | -106.4846
16.756 | -106.4957
16.7596 | -106.5035
16.7632 | -106.5122
16.7668 | -106.524
16.7704 | -106.5359
16.774 | -106.5483
16.7776 | -106.5597
16.7812 | -106.5736
16.7849 | -106.5853
16.7885 | -106.5944
16.7921 | -106.6046
16.7957 | -106.617
16.7993 | -106.6302
16.8029 | -106.6409
16.8065 | -106.6526
16.8102 | -106.6654
16.8138 | -106.676
16.8174 | -106.6856
16.821 | -106.6968
16.8246 | -106.7081
16.8282 | -106.7226
16.8318 | -106.7349
16.8354 | -106.7451
16.8391 | -106.7592
16.8427 | -106.7728
16.8463 | -106.7828
16.8499 | -106.7929
16.8535 | -106.8027
16.8571 | -106.8145
16.8607 | -106.8274
16.8644 | -106.8401
16.868 | -106.8506
16.8716 | -106.8626
16.8752 | -106.8744
16.8788 | -106.8869
16.8824 | -106.9005
16.886 | -106.9147
16.8897 | -106.9264
16.8933 | -106.937
16.8969 | -106.9456
16.9005 | -106.955
16.9041 | -106.9663
16.9077 | -106.9753
16.9113 | -106.9845
16.9149 | -106.9941
16.9186 | -107.0003
16.9222 | -107.0043
16.9258 | -107.0051
16.9294 | -107.0068
16.933 | -107.0085
16.9366 | -107.009
16.9402 | -107.0072
16.9439 | -107.002
16.9475 | -106.9951
16.9511 | -106.9882
16.9547 | -106.9785
16.9583 | -106.9677
16.9619 | -106.9572
16.9655 | -106.9441
16.9691 | -106.9321
16.9728 | -106.9186
16.9764 | -106.903
16.98 | -106.887
16.9836 | -106.8687
16.9872 | -106.851
16.9908 | -106.8346
16.9944 | -106.8199
16.9981 | -106.8047
17.0017 | -106.7916
17.0053 | -106.7762
17.0089 | -106.7595
17.0125 | -106.7418
17.0161 | -106.7254
17.0197 | -106.7113
17.0234 | -106.6964
17.027 | -106.681
17.0306 | -106.6685
17.0342 | -106.6575
17.0378 | -106.6468
17.0414 | -106.6393
17.045 | -106.6326
17.0486 | -106.6265
17.0523 | -106.6205
17.0559 | -106.6179
17.0595 | -106.6197
17.0631 | -106.6188
17.0667 | -106.6199
17.0703 | -106.6221
17.0739 | -106.6249
17.0776 | -106.6284
17.0812 | -106.6317
17.0848 | -106.6357
17.0884 | -106.6389
17.092 | -106.6464
17.0956 | -106.6538
17.0992 | -106.6593
17.1029 | -106.6645
17.1065 | -106.6698
17.1101 | -106.6742
17.1137 | -106.678
17.1173 | -106.6837
17.1209 | -106.6889
17.1245 | -106.6923
17.1281 | -106.6949
17.1318 | -106.6979
17.1354 | -106.6998
17.139 | -106.7018
17.1426 | -106.7016
17.1462 | -106.6983
17.1498 | -106.6941
17.1534 | -106.6874
17.1571 | -106.6843
17.1607 | -106.6817
17.1643 | -106.678
17.1679 | -106.6765
17.1715 | -106.6716
17.1751 | -106.6649
17.1787 | -106.66
17.1823 | -106.6562
17.186 | -106.6524
17.1896 | -106.6487
17.1932 | -106.6454
17.1968 | -106.6415
17.2004 | -106.6419
17.204 | -106.6399
17.2076 | -106.6381
17.2113 | -106.6371
17.2149 | -106.6345
17.2185 | -106.6334
17.2221 | -106.6329
17.2257 | -106.6322
17.2293 | -106.6321
17.2329 | -106.6325
17.2366 | -106.6337
17.2402 | -106.6351
17.2438 | -106.6399
17.2474 | -106.6433
17.251 | -106.6502
17.2546 | -106.658
17.2582 | -106.6668
17.2618 | -106.6749
17.2655 | -106.6839
17.2691 | -106.693
17.2727 | -106.7046
17.2763 | -106.716
17.2799 | -106.7279
17.2835 | -106.7408
17.2871 | -106.75
17.2908 | -106.7592
17.2944 | -106.771
17.298 | -106.7826
17.3016 | -106.7912
17.3052 | -106.7986
17.3088 | -106.8099
17.3124 | -106.8186
17.316 | -106.8243
17.3197 | -106.8315
17.3233 | -106.8377
17.3269 | -106.8388
17.3305 | -106.8384
17.3341 | -106.839
17.3377 | -106.8365
17.3413 | -106.8322
17.345 | -106.8258
17.3486 | -106.8186
17.3522 | -106.8117
17.3558 | -106.8042
17.3594 | -106.7928
17.363 | -106.7794
17.3666 | -106.7649
17.3703 | -106.7509
17.3739 | -106.7375
17.3775 | -106.7216
17.3811 | -106.7051
17.3847 | -106.6885
17.3883 | -106.6715
17.3919 | -106.6573
17.3955 | -106.6442
17.3992 | -106.6307
17.4028 | -106.613
17.4064 | -106.5955
17.41 | -106.5817
17.4136 | -106.5678
17.4172 | -106.5578
17.4208 | -106.5495
17.4245 | -106.5373
17.4281 | -106.5276
17.4317 | -106.5187
17.4353 | -106.5118
17.4389 | -106.5055
17.4425 | -106.5009
17.4461 | -106.4952
17.4497 | -106.4896
17.4534 | -106.4858
17.457 | -106.4849
17.4606 | -106.4855
17.4642 | -106.4878
17.4678 | -106.4923
17.4714 | -106.4982
17.475 | -106.5034
17.4787 | -106.5096
17.4823 | -106.5161
17.4859 | -106.5226
17.4895 | -106.5309
17.4931 | -106.5375
17.4967 | -106.5442
17.5003 | -106.5502
17.504 | -106.5558
17.5076 | -106.5628
17.5112 | -106.5726
17.5148 | -106.5802
17.5184 | -106.589
17.522 | -106.5975
17.5256 | -106.6051
17.5292 | -106.6115
17.5329 | -106.6173
17.5365 | -106.626
17.5401 | -106.6323
17.5437 | -106.6374
17.5473 | -106.6439
17.5509 | -106.6528
17.5545 | -106.6602
17.5582 | -106.6637
17.5618 | -106.6671
17.5654 | -106.6703
17.569 | -106.6746
17.5726 | -106.6791
17.5762 | -106.685
17.5798 | -106.6915
17.5834 | -106.6964
17.5871 | -106.7007
17.5907 | -106.7045
17.5943 | -106.7093
17.5979 | -106.7126
17.6015 | -106.7144
17.6051 | -106.7171
17.6087 | -106.7199
17.6124 | -106.7197
17.616 | -106.7217
17.6196 | -106.7236
17.6232 | -106.7229
17.6268 | -106.7217
17.6304 | -106.7199
17.634 | -106.7182
17.6377 | -106.716
17.6413 | -106.7146
17.6449 | -106.7132
17.6485 | -106.7123
17.6521 | -106.71
17.6557 | -106.7072
17.6593 | -106.7044
17.6629 | -106.7025
17.6666 | -106.6994
17.6702 | -106.6993
17.6738 | -106.7022
17.6774 | -106.7048
17.681 | -106.709
17.6846 | -106.7113
17.6882 | -106.7136
17.6919 | -106.7161
17.6955 | -106.7193
17.6991 | -106.723
17.7027 | -106.7248
17.7063 | -106.7253
17.7099 | -106.7261
17.7135 | -106.7263
17.7172 | -106.7247
17.7208 | -106.7235
17.7244 | -106.7279
17.728 | -106.7316
17.7316 | -106.7339
17.7352 | -106.7356
17.7388 | -106.7388
17.7424 | -106.7425
17.7461 | -106.7454
17.7497 | -106.7501
17.7533 | -106.7544
17.7569 | -106.759
17.7605 | -106.763
17.7641 | -106.7625
17.7677 | -106.7647
17.7714 | -106.7666
17.775 | -106.7682
17.7786 | -106.7725
17.7822 | -106.7755
17.7858 | -106.7788
17.7894 | -106.7796
17.793 | -106.7813
17.7966 | -106.7832
17.8003 | -106.7839
17.8039 | -106.7843
17.8075 | -106.7856
17.8111 | -106.7888
17.8147 | -106.7923
17.8183 | -106.796
17.8219 | -106.796
17.8256 | -106.7942
17.8292 | -106.795
17.8328 | -106.7947
17.8364 | -106.794
17.84 | -106.7933
17.8436 | -106.7954
17.8472 | -106.7954
17.8509 | -106.7933
17.8545 | -106.7947
17.8581 | -106.7953
17.8617 | -106.7939
17.8653 | -106.7931
17.8689 | -106.7919
17.8725 | -106.7904
17.8761 | -106.79
17.8798 | -106.7909
17.8834 | -106.7915
17.887 | -106.7887
17.8906 | -106.7853
17.8942 | -106.7788
17.8978 | -106.7717
17.9014 | -106.765
17.9051 | -106.7602
17.9087 | -106.7548
17.9123 | -106.7478
17.9159 | -106.7384
17.9195 | -106.7306
17.9231 | -106.72
17.9267 | -106.7095
17.9303 | -106.6995
17.934 | -106.6903
17.9376 | -106.6835
17.9412 | -106.6763
17.9448 | -106.6693
17.9484 | -106.6619
17.952 | -106.6561
17.9556 | -106.6512
17.9593 | -106.6441
17.9629 | -106.6358
17.9665 | -106.6311
17.9701 | -106.6255
17.9737 | -106.6173
17.9773 | -106.6102
17.9809 | -106.6026
17.9846 | -106.5955
17.9882 | -106.5904
17.9918 | -106.5853
17.9954 | -106.58
17.999 | -106.5726
18.0026 | -106.5677
18.0062 | -106.5634
18.0098 | -106.5612
18.0135 | -106.5605
18.0171 | -106.5602
18.0207 | -106.5625
18.0243 | -106.5629
18.0279 | -106.5626
18.0315 | -106.5635
18.0351 | -106.5628
18.0388 | -106.5639
18.0424 | -106.5643
18.046 | -106.5663
18.0496 | -106.5699
18.0532 | -106.5724
18.0568 | -106.5749
18.0604 | -106.5771
18.0641 | -106.5766
18.0677 | -106.5769
18.0713 | -106.5787
18.0749 | -106.578
18.0785 | -106.5784
18.0821 | -106.5809
18.0857 | -106.5804
18.0893 | -106.5805
18.093 | -106.581
18.0966 | -106.581
18.1002 | -106.5782
18.1038 | -106.5755
18.1074 | -106.5727
18.111 | -106.5689
18.1146 | -106.5647
18.1183 | -106.5611
18.1219 | -106.5582
18.1255 | -106.5543
18.1291 | -106.5492
18.1327 | -106.544
18.1363 | -106.5379
18.1399 | -106.5339
18.1435 | -106.5332
18.1472 | -106.5333
18.1508 | -106.5334
18.1544 | -106.5325
18.158 | -106.533
18.1616 | -106.5334
18.1652 | -106.5341
18.1688 | -106.5367
18.1725 | -106.5413
18.1761 | -106.5447
18.1797 | -106.5502
18.1833 | -106.5566
18.1869 | -106.5647
18.1905 | -106.5734
18.1941 | -106.5842
18.1978 | -106.5969
18.2014 | -106.6091
18.205 | -106.6225
18.2086 | -106.6345
18.2122 | -106.6472
18.2158 | -106.6636
18.2194 | -106.6788
18.223 | -106.6913
18.2267 | -106.7074
18.2303 | -106.7233
18.2339 | -106.7377
18.2375 | -106.7509
18.2411 | -106.7652
18.2447 | -106.7793
18.2483 | -106.793
18.252 | -106.8078
18.2556 | -106.8218
18.2592 | -106.8368
18.2628 | -106.848
18.2664 | -106.8572
18.27 | -106.8664
18.2736 | -106.8754
18.2772 | -106.884
18.2809 | -106.8924
18.2845 | -106.9007
18.2881 | -106.9087
18.2917 | -106.9123
18.2953 | -106.9158
18.2989 | -106.9194
18.3025 | -106.923
18.3062 | -106.9236
18.3098 | -106.923
18.3134 | -106.9233
18.317 | -106.9245
18.3206 | -106.922
18.3242 | -106.919
18.3278 | -106.9133
18.3315 | -106.9044
18.3351 | -106.897
18.3387 | -106.8887
18.3423 | -106.8797
18.3459 | -106.8716
18.3495 | -106.8621
18.3531 | -106.8518
18.3567 | -106.8394
18.3604 | -106.8288
18.364 | -106.821
18.3676 | -106.8102
18.3712 | -106.7954
18.3748 | -106.7806
18.3784 | -106.7662
18.382 | -106.7525
18.3857 | -106.7361
18.3893 | -106.7213
18.3929 | -106.7062
18.3965 | -106.6917
18.4001 | -106.6744
18.4037 | -106.6569
18.4073 | -106.6417
18.4109 | -106.6289
18.4146 | -106.6167
18.4182 | -106.6027
18.4218 | -106.5904
18.4254 | -106.5792
18.429 | -106.568
18.4326 | -106.5571
18.4362 | -106.5467
18.4399 | -106.5383
18.4435 | -106.5315
18.4471 | -106.5231
18.4507 | -106.5159
18.4543 | -106.5091
18.4579 | -106.5017
18.4615 | -106.4932
18.4652 | -106.4873
18.4688 | -106.4835
18.4724 | -106.4797
18.476 | -106.476
18.4796 | -106.4737
18.4832 | -106.4729
18.4868 | -106.4716
18.4904 | -106.4715
18.4941 | -106.4692
18.4977 | -106.469
18.5013 | -106.4673
18.5049 | -106.4633
18.5085 | -106.4596
18.5121 | -106.4579
18.5157 | -106.4561
18.5194 | -106.4544
18.523 | -106.4517
18.5266 | -106.4522
18.5302 | -106.4549
18.5338 | -106.4556
18.5374 | -106.4553
18.541 | -106.4566
18.5447 | -106.458
18.5483 | -106.4614
18.5519 | -106.4653
18.5555 | -106.4706
18.5591 | -106.4748
18.5627 | -106.4767
18.5663 | -106.4809
18.5699 | -106.4854
18.5736 | -106.4895
18.5772 | -106.4926
18.5808 | -106.4962
18.5844 | -106.5019
18.588 | -106.5087
18.5916 | -106.5172
18.5952 | -106.5262
18.5989 | -106.534
18.6025 | -106.5408
18.6061 | -106.5461
18.6097 | -106.5492
18.6133 | -106.5551
18.6169 | -106.5619
18.6205 | -106.5703
18.6241 | -106.5771
18.6278 | -106.5823
18.6314 | -106.5896
18.635 | -106.5977
18.6386 | -106.6076
18.6422 | -106.6168
18.6458 | -106.6216
18.6494 | -106.63
18.6531 | -106.6377
18.6567 | -106.6453
18.6603 | -106.6514
18.6639 | -106.6575
18.6675 | -106.6641
18.6711 | -106.6695
18.6747 | -106.6732
18.6783 | -106.6757
18.682 | -106.6761
18.6856 | -106.6745
18.6892 | -106.6725
18.6928 | -106.6672
18.6964 | -106.6624
18.7 | -106.6579
18.7036 | -106.6501
18.7073 | -106.6393
18.7109 | -106.6274
18.7145 | -106.6151
18.7181 | -106.6031
18.7217 | -106.592
18.7253 | -106.5807
18.7289 | -106.565
18.7326 | -106.5471
18.7362 | -106.53
18.7398 | -106.5116
18.7434 | -106.4917
18.747 | -106.4692
18.7506 | -106.4459
18.7542 | -106.4253
18.7578 | -106.4033
18.7615 | -106.3797
18.7651 | -106.3551
18.7687 | -106.3319
18.7723 | -106.3103
18.7759 | -106.2892
18.7795 | -106.267
18.7831 | -106.246
18.7868 | -106.2265
18.7904 | -106.2081
18.794 | -106.1901
18.7976 | -106.1699
18.8012 | -106.1515
18.8048 | -106.1364
18.8084 | -106.1214
18.8121 | -106.1062
18.8157 | -106.0908
18.8193 | -106.0758
18.8229 | -106.0642
18.8265 | -106.0542
18.8301 | -106.0465
18.8337 | -106.0391
18.8373 | -106.034
18.841 | -106.0308
18.8446 | -106.0266
18.8482 | -106.0218
18.8518 | -106.0192
18.8554 | -106.0178
18.859 | -106.0179
18.8626 | -106.0205
18.8663 | -106.024
18.8699 | -106.0281
18.8735 | -106.0333
18.8771 | -106.0408
18.8807 | -106.0468
18.8843 | -106.0537
18.8879 | -106.0587
18.8915 | -106.0654
18.8952 | -106.0737
18.8988 | -106.0827
18.9024 | -106.0936
18.906 | -106.1025
18.9096 | -106.1106
18.9132 | -106.1212
18.9168 | -106.1286
18.9205 | -106.1364
18.9241 | -106.1444
18.9277 | -106.1523
18.9313 | -106.1597
18.9349 | -106.1669
18.9385 | -106.1744
18.9421 | -106.1811
18.9458 | -106.1894
18.9494 | -106.1994
18.953 | -106.2089
18.9566 | -106.2179
18.9602 | -106.2266
18.9638 | -106.2333
18.9674 | -106.2385
18.971 | -106.246
18.9747 | -106.2532
18.9783 | -106.2596
18.9819 | -106.2637
18.9855 | -106.2668
18.9891 | -106.2692
18.9927 | -106.2731
18.9963 | -106.2776
19 | -106.2805
19.0036 | -106.2836
19.0072 | -106.2853
19.0108 | -106.287
19.0144 | -106.2895
19.018 | -106.2929
19.0216 | -106.2953
19.0252 | -106.2987
19.0289 | -106.2998
19.0325 | -106.3002
19.0361 | -106.3018
19.0397 | -106.3049
19.0433 | -106.3052
19.0469 | -106.3047
19.0505 | -106.3033
19.0542 | -106.3013
19.0578 | -106.2995
19.0614 | -106.2986
19.065 | -106.2961
19.0686 | -106.2949
19.0722 | -106.2934
19.0758 | -106.2902
19.0795 | -106.2857
19.0831 | -106.2835
19.0867 | -106.2794
19.0903 | -106.2771
19.0939 | -106.2739
19.0975 | -106.2681
19.1011 | -106.2592
19.1047 | -106.2482
19.1084 | -106.2357
19.112 | -106.224
19.1156 | -106.2139
19.1192 | -106.2028
19.1228 | -106.1889
19.1264 | -106.1722
19.13 | -106.1539
19.1337 | -106.1354
19.1373 | -106.117
19.1409 | -106.1
19.1445 | -106.0827
19.1481 | -106.0629
19.1517 | -106.042
19.1553 | -106.0205
19.159 | -106.0009
19.1626 | -105.9817
19.1662 | -105.9603
19.1698 | -105.9392
19.1734 | -105.9177
19.177 | -105.8967
19.1806 | -105.8762
19.1842 | -105.8569
19.1879 | -105.8408
19.1915 | -105.8234
19.1951 | -105.8078
19.1987 | -105.7936
19.2023 | -105.781
19.2059 | -105.7684
19.2095 | -105.7586
19.2132 | -105.7505
19.2168 | -105.7427
19.2204 | -105.7366
19.224 | -105.7334
19.2276 | -105.7311
19.2312 | -105.7311
19.2348 | -105.7318
19.2384 | -105.7346
19.2421 | -105.7376
19.2457 | -105.742
19.2493 | -105.749
19.2529 | -105.7573
19.2565 | -105.7658
19.2601 | -105.7763
19.2637 | -105.7865
19.2674 | -105.7958
19.271 | -105.8069
19.2746 | -105.8162
19.2782 | -105.8263
19.2818 | -105.8372
19.2854 | -105.8444
19.289 | -105.8534
19.2927 | -105.8616
19.2963 | -105.869
19.2999 | -105.876
19.3035 | -105.8839
19.3071 | -105.8917
19.3107 | -105.896
19.3143 | -105.9008
19.3179 | -105.9071
19.3216 | -105.9131
19.3252 | -105.9178
19.3288 | -105.9204
19.3324 | -105.92
19.336 | -105.9197
19.3396 | -105.9198
19.3432 | -105.9197
19.3469 | -105.9163
19.3505 | -105.9137
19.3541 | -105.9133
19.3577 | -105.9117
19.3613 | -105.9094
19.3649 | -105.9071
19.3685 | -105.9027
19.3721 | -105.9014
19.3758 | -105.9032
19.3794 | -105.9046
19.383 | -105.9064
19.3866 | -105.9062
19.3902 | -105.9052
19.3938 | -105.9083
19.3974 | -105.9135
19.4011 | -105.9196
19.4047 | -105.9261
19.4083 | -105.9327
19.4119 | -105.9399
19.4155 | -105.9461
19.4191 | -105.953
19.4227 | -105.9621
19.4264 | -105.9724
19.43 | -105.9831
19.4336 | -105.9942
19.4372 | -106.0047
19.4408 | -106.0167
19.4444 | -106.028
19.448 | -106.0385
19.4516 | -106.049
19.4553 | -106.0623
19.4589 | -106.0754
19.4625 | -106.0874
19.4661 | -106.0997
19.4697 | -106.1101
19.4733 | -106.1224
19.4769 | -106.1323
19.4806 | -106.1399
19.4842 | -106.1448
19.4878 | -106.1486
19.4914 | -106.1513
19.495 | -106.1526
19.4986 | -106.1534
19.5022 | -106.1534
19.5058 | -106.1544
19.5095 | -106.1565
19.5131 | -106.1542
19.5167 | -106.1518
19.5203 | -106.1492
19.5239 | -106.1451
19.5275 | -106.1409
19.5311 | -106.1361
19.5348 | -106.1303
19.5384 | -106.1245
19.542 | -106.1189
19.5456 | -106.1131
19.5492 | -106.1086
19.5528 | -106.1034
19.5564 | -106.0971
19.5601 | -106.09
19.5637 | -106.083
19.5673 | -106.0758
19.5709 | -106.0684
19.5745 | -106.0616
19.5781 | -106.0554
19.5817 | -106.0498
19.5853 | -106.0449
19.589 | -106.0406
19.5926 | -106.0347
19.5962 | -106.0296
19.5998 | -106.0255
19.6034 | -106.0225
19.607 | -106.0224
19.6106 | -106.0237
19.6143 | -106.0224
19.6179 | -106.0205
19.6215 | -106.021
19.6251 | -106.0252
19.6287 | -106.0283
19.6323 | -106.0296
19.6359 | -106.0325
19.6395 | -106.0383
19.6432 | -106.0426
19.6468 | -106.0472
19.6504 | -106.051
19.654 | -106.0553
19.6576 | -106.0591
19.6612 | -106.0614
19.6648 | -106.0659
19.6685 | -106.0705
19.6721 | -106.0713
19.6757 | -106.0727
19.6793 | -106.0765
19.6829 | -106.0787
19.6865 | -106.0816
19.6901 | -106.0843
19.6938 | -106.088
19.6974 | -106.0928
19.701 | -106.0983
19.7046 | -106.1018
19.7082 | -106.1062
19.7118 | -106.1104
19.7154 | -106.1122
19.719 | -106.1137
19.7227 | -106.1157
19.7263 | -106.1191
19.7299 | -106.1229
19.7335 | -106.1242
19.7371 | -106.1236
19.7407 | -106.1235
19.7443 | -106.1262
19.748 | -106.1288
19.7516 | -106.1294
19.7552 | -106.1318
19.7588 | -106.1343
19.7624 | -106.1375
19.766 | -106.1432
19.7696 | -106.1475
19.7733 | -106.1508
19.7769 | -106.1519
19.7805 | -106.1527
19.7841 | -106.1534
19.7877 | -106.1536
19.7913 | -106.1553
19.7949 | -106.1565
19.7985 | -106.1557
19.8022 | -106.1568
19.8058 | -106.158
19.8094 | -106.1572
19.813 | -106.1589
19.8166 | -106.1599
19.8202 | -106.1613
19.8238 | -106.1641
19.8275 | -106.1656
19.8311 | -106.1671
19.8347 | -106.1695
19.8383 | -106.1725
19.8419 | -106.1766
19.8455 | -106.181
19.8491 | -106.1862
19.8527 | -106.1913
19.8564 | -106.1947
19.86 | -106.1971
19.8636 | -106.2
19.8672 | -106.2026
19.8708 | -106.2052
19.8744 | -106.2093
19.878 | -106.2118
19.8817 | -106.2151
19.8853 | -106.2201
19.8889 | -106.2229
19.8925 | -106.226
19.8961 | -106.2305
19.8997 | -106.2349
19.9033 | -106.2368
19.907 | -106.2376
19.9106 | -106.2379
19.9142 | -106.2395
19.9178 | -106.2421
19.9214 | -106.2437
19.925 | -106.2467
19.9286 | -106.2501
19.9322 | -106.2537
19.9359 | -106.2567
19.9395 | -106.2616
19.9431 | -106.269
19.9467 | -106.2734
19.9503 | -106.2775
19.9539 | -106.2828
19.9575 | -106.2875
19.9612 | -106.2903
19.9648 | -106.2943
19.9684 | -106.2986
19.972 | -106.3001
19.9756 | -106.3032
19.9792 | -106.3065
19.9828 | -106.3072
19.9864 | -106.3091
19.9901 | -106.3112
19.9937 | -106.3132
19.9973 | -106.3161
20.0009 | -106.3202
20.0045 | -106.3243
20.0081 | -106.3269
20.0117 | -106.3306
20.0154 | -106.3334
20.019 | -106.337
20.0226 | -106.3399
20.0262 | -106.3384
20.0298 | -106.3377
20.0334 | -106.3385
20.037 | -106.3401
20.0407 | -106.3437
20.0443 | -106.3444
20.0479 | -106.3416
20.0515 | -106.3418
20.0551 | -106.3434
20.0587 | -106.3446
20.0623 | -106.3458
20.0659 | -106.3467
20.0696 | -106.3459
20.0732 | -106.3446
20.0768 | -106.3452
20.0804 | -106.3468
20.084 | -106.3472
20.0876 | -106.3458
20.0912 | -106.3456
20.0949 | -106.3465
20.0985 | -106.3454
20.1021 | -106.3445
20.1057 | -106.3446
20.1093 | -106.3414
20.1129 | -106.3389
20.1165 | -106.3361
20.1201 | -106.3337
20.1238 | -106.3328
20.1274 | -106.3304
20.131 | -106.3285
20.1346 | -106.3281
20.1382 | -106.3266
20.1418 | -106.3272
20.1454 | -106.3284
20.1491 | -106.3275
20.1527 | -106.3257
20.1563 | -106.3236
20.1599 | -106.3211
20.1635 | -106.3214
20.1671 | -106.3219
20.1707 | -106.3197
20.1744 | -106.3182
20.178 | -106.3162
20.1816 | -106.3165
20.1852 | -106.3166
20.1888 | -106.317
20.1924 | -106.318
20.196 | -106.3186
20.1996 | -106.3193
20.2033 | -106.3194
20.2069 | -106.3204
20.2105 | -106.3215
20.2141 | -106.321
20.2177 | -106.319
20.2213 | -106.318
20.2249 | -106.3177
20.2286 | -106.3171
20.2322 | -106.3163
20.2358 | -106.3157
20.2394 | -106.3157
20.243 | -106.3165
20.2466 | -106.3175
20.2502 | -106.3178
20.2539 | -106.3163
20.2575 | -106.3159
20.2611 | -106.3131
20.2647 | -106.3098
20.2683 | -106.3098
20.2719 | -106.311
20.2755 | -106.3076
20.2791 | -106.3043
20.2828 | -106.301
20.2864 | -106.2948
20.29 | -106.2885
20.2936 | -106.283
20.2972 | -106.2771
20.3008 | -106.273
20.3044 | -106.2704
20.3081 | -106.2657
20.3117 | -106.2615
20.3153 | -106.2584
20.3189 | -106.2545
20.3225 | -106.2496
20.3261 | -106.2403
20.3297 | -106.2226
20.3333 | -106.2154
20.337 | -106.2098
20.3406 | -106.2025
20.3442 | -106.1933
20.3478 | -106.1826
20.3514 | -106.1729
20.355 | -106.1629
20.3586 | -106.1533
20.3623 | -106.1447
20.3659 | -106.1335
20.3695 | -106.1223
20.3731 | -106.1091
20.3767 | -106.0971
20.3803 | -106.0858
20.3839 | -106.0722
20.3876 | -106.0564
20.3912 | -106.042
20.3948 | -106.0267
20.3984 | -106.0144
20.402 | -106.0016
20.4056 | -105.9878
20.4092 | -105.9737
20.4128 | -105.9593
20.4165 | -105.9448
20.4201 | -105.9282
20.4237 | -105.9132
20.4273 | -105.8997
20.4309 | -105.8867
20.4345 | -105.8757
20.4381 | -105.866
20.4418 | -105.8564
20.4454 | -105.8477
20.449 | -105.8387
20.4526 | -105.8305
20.4562 | -105.8229
20.4598 | -105.8165
20.4634 | -105.8088
20.467 | -105.8027
20.4707 | -105.7974
20.4743 | -105.7963
20.4779 | -105.794
20.4815 | -105.7898
20.4851 | -105.7865
20.4887 | -105.7833
20.4923 | -105.7809
20.496 | -105.7813
20.4996 | -105.7817
20.5032 | -105.7843
20.5068 | -105.7879
20.5104 | -105.7897
20.514 | -105.7909
20.5176 | -105.7926
20.5213 | -105.7969
20.5249 | -105.8013
20.5285 | -105.8052
20.5321 | -105.8077
20.5357 | -105.809
20.5393 | -105.8109
20.5429 | -105.8141
20.5465 | -105.8181
20.5502 | -105.8219
20.5538 | -105.825
20.5574 | -105.8267
20.561 | -105.8273
20.5646 | -105.8311
20.5682 | -105.8346
20.5718 | -105.8364
20.5755 | -105.8386
20.5791 | -105.8403
20.5827 | -105.8404
20.5863 | -105.8399
20.5899 | -105.8386
20.5935 | -105.8364
20.5971 | -105.8351
20.6008 | -105.8355
20.6044 | -105.8343
20.608 | -105.8306
20.6116 | -105.8274
20.6152 | -105.8284
20.6188 | -105.8323
20.6224 | -105.832
20.626 | -105.8303
20.6297 | -105.8281
20.6333 | -105.8255
20.6369 | -105.8233
20.6405 | -105.8201
20.6441 | -105.8161
20.6477 | -105.8138
20.6513 | -105.8129
20.655 | -105.8115
20.6586 | -105.8103
20.6622 | -105.8104
20.6658 | -105.8122
20.6694 | -105.8148
20.673 | -105.8165
20.6766 | -105.8191
20.6802 | -105.8223
20.6839 | -105.8253
20.6875 | -105.8291
20.6911 | -105.8325
20.6947 | -105.8363
20.6983 | -105.8379
20.7019 | -105.8383
20.7055 | -105.8413
20.7092 | -105.845
20.7128 | -105.848
20.7164 | -105.8504
20.72 | -105.8536
20.7236 | -105.8572
20.7272 | -105.8594
20.7308 | -105.8618
20.7345 | -105.8645
20.7381 | -105.8665
20.7417 | -105.8698
20.7453 | -105.8696
20.7489 | -105.8697
20.7525 | -105.8706
20.7561 | -105.8737
20.7597 | -105.8766
20.7634 | -105.8773
20.767 | -105.8759
20.7706 | -105.8753
20.7742 | -105.8754
20.7778 | -105.8742
20.7814 | -105.8735
20.785 | -105.8717
20.7887 | -105.8658
20.7923 | -105.8614
20.7959 | -105.8586
20.7995 | -105.8571
20.8031 | -105.8557
20.8067 | -105.8544
20.8103 | -105.8541
20.8139 | -105.8501
20.8176 | -105.8482
20.8212 | -105.8459
20.8248 | -105.8441
20.8284 | -105.8425
20.832 | -105.8415
20.8356 | -105.8395
20.8392 | -105.8374
20.8429 | -105.8361
20.8465 | -105.8338
20.8501 | -105.8338
20.8537 | -105.8334
20.8573 | -105.833
20.8609 | -105.8318
20.8645 | -105.8314
20.8682 | -105.8322
20.8718 | -105.8339
20.8754 | -105.8359
20.879 | -105.8391
20.8826 | -105.8425
20.8862 | -105.8463
20.8898 | -105.8502
20.8934 | -105.8546
20.8971 | -105.8588
20.9007 | -105.8631
20.9043 | -105.8701
20.9079 | -105.8777
20.9115 | -105.8839
20.9151 | -105.8887
20.9187 | -105.896
20.9224 | -105.9015
20.926 | -105.9091
20.9296 | -105.9184
20.9332 | -105.9288
20.9368 | -105.9405
20.9404 | -105.9503
20.944 | -105.9591
20.9476 | -105.9687
20.9513 | -105.9784
20.9549 | -105.9883
20.9585 | -105.9979
20.9621 | -106.0075
20.9657 | -106.0167
20.9693 | -106.026
20.9729 | -106.0354
20.9766 | -106.0474
20.9802 | -106.0583
20.9838 | -106.0657
20.9874 | -106.0737
20.991 | -106.0832
20.9946 | -106.0962
20.9982 | -106.1076
21.0019 | -106.1158
21.0055 | -106.1236
21.0091 | -106.1346
21.0127 | -106.1441
21.0163 | -106.1527
21.0199 | -106.1574
21.0235 | -106.1621
21.0271 | -106.1679
21.0308 | -106.1743
21.0344 | -106.181
21.038 | -106.1867
21.0416 | -106.1923
21.0452 | -106.1983
21.0488 | -106.205
21.0524 | -106.2113
21.0561 | -106.2152
21.0597 | -106.2177
21.0633 | -106.2209
21.0669 | -106.2261
21.0705 | -106.2313
21.0741 | -106.2359
21.0777 | -106.2391
21.0814 | -106.2414
21.085 | -106.2453
21.0886 | -106.2476
21.0922 | -106.2485
21.0958 | -106.2467
21.0994 | -106.2459
21.103 | -106.2462
21.1066 | -106.2445
21.1103 | -106.2464
21.1139 | -106.2455
21.1175 | -106.2449
21.1211 | -106.2443
21.1247 | -106.2453
21.1283 | -106.246
21.1319 | -106.247
21.1356 | -106.248
21.1392 | -106.2465
21.1428 | -106.2457
21.1464 | -106.2467
21.15 | -106.2479
21.1536 | -106.2479
21.1572 | -106.2484
21.1608 | -106.2474
21.1645 | -106.2458
21.1681 | -106.2455
21.1717 | -106.2438
21.1753 | -106.2419
21.1789 | -106.2413
21.1825 | -106.2416
21.1861 | -106.24
21.1898 | -106.2399
21.1934 | -106.2411
21.197 | -106.2428
21.2006 | -106.2434
21.2042 | -106.2436
21.2078 | -106.242
21.2114 | -106.2399
21.2151 | -106.2398
21.2187 | -106.239
21.2223 | -106.239
21.2259 | -106.2378
21.2295 | -106.2349
21.2331 | -106.2335
21.2367 | -106.2331
21.2403 | -106.232
21.244 | -106.2299
21.2476 | -106.2277
21.2512 | -106.2241
21.2548 | -106.2207
21.2584 | -106.2184
21.262 | -106.2163
21.2656 | -106.2119
21.2693 | -106.2102
21.2729 | -106.2089
21.2765 | -106.2057
21.2801 | -106.2034
21.2837 | -106.2044
21.2873 | -106.2032
21.2909 | -106.2012
21.2945 | -106.1976
21.2982 | -106.198
21.3018 | -106.1976
21.3054 | -106.1972
21.309 | -106.1954
21.3126 | -106.1935
21.3162 | -106.1924
21.3198 | -106.194
21.3235 | -106.1953
21.3271 | -106.1961
21.3307 | -106.1959
21.3343 | -106.1969
21.3379 | -106.1984
21.3415 | -106.201
21.3451 | -106.2024
21.3488 | -106.2018
21.3524 | -106.2015
21.356 | -106.2043
21.3596 | -106.2087
21.3632 | -106.2113
21.3668 | -106.214
21.3704 | -106.2169
21.374 | -106.2222
21.3777 | -106.2263
21.3813 | -106.2296
21.3849 | -106.2332
21.3885 | -106.236
21.3921 | -106.2393
21.3957 | -106.2413
21.3993 | -106.2432
21.403 | -106.246
21.4066 | -106.2475
21.4102 | -106.2479
21.4138 | -106.2484
21.4174 | -106.2483
21.421 | -106.25
21.4246 | -106.2486
21.4282 | -106.2455
21.4319 | -106.2429
21.4355 | -106.2385
21.4391 | -106.2351
21.4427 | -106.2325
21.4463 | -106.2277
21.4499 | -106.2173
21.4535 | -106.2077
21.4572 | -106.2001
21.4608 | -106.1907
21.4644 | -106.1828
21.468 | -106.1747
21.4716 | -106.166
21.4752 | -106.1554
21.4788 | -106.1427
21.4825 | -106.1295
21.4861 | -106.1156
21.4897 | -106.1035
21.4933 | -106.0898
21.4969 | -106.0749
21.5005 | -106.0603
21.5041 | -106.047
21.5077 | -106.034
21.5114 | -106.0225
21.515 | -106.0115
21.5186 | -105.9988
21.5222 | -105.9871
21.5258 | -105.9743
21.5294 | -105.9622
21.533 | -105.9521
21.5367 | -105.9406
21.5403 | -105.9299
21.5439 | -105.9221
21.5475 | -105.9153
21.5511 | -105.9063
21.5547 | -105.8986
21.5583 | -105.8906
21.5619 | -105.882
21.5656 | -105.8754
21.5692 | -105.8726
21.5728 | -105.8675
21.5764 | -105.863
21.58 | -105.86
21.5836 | -105.8568
21.5872 | -105.8539
21.5909 | -105.8523
21.5945 | -105.8513
21.5981 | -105.8508
21.6017 | -105.8496
21.6053 | -105.8495
21.6089 | -105.8513
21.6125 | -105.8542
21.6162 | -105.8576
21.6198 | -105.8591
21.6234 | -105.862
21.627 | -105.866
21.6306 | -105.8698
21.6342 | -105.8746
21.6378 | -105.8783
21.6414 | -105.8817
21.6451 | -105.8857
21.6487 | -105.8892
21.6523 | -105.8957
21.6559 | -105.9022
21.6595 | -105.9078
21.6631 | -105.9138
21.6667 | -105.9204
21.6704 | -105.9264
21.674 | -105.9333
21.6776 | -105.9373
21.6812 | -105.9416
21.6848 | -105.9467
21.6884 | -105.9535
21.692 | -105.9602
21.6957 | -105.9664
21.6993 | -105.9694
21.7029 | -105.9742
21.7065 | -105.98
21.7101 | -105.9824
21.7137 | -105.9826
21.7173 | -105.987
21.7209 | -105.9907
21.7246 | -105.9925
21.7282 | -105.9906
21.7318 | -105.9921
21.7354 | -105.9932
21.739 | -105.9912
21.7426 | -105.9908
21.7462 | -105.9892
21.7499 | -105.9849
21.7535 | -105.9805
21.7571 | -105.9792
21.7607 | -105.9752
21.7643 | -105.9688
21.7679 | -105.9628
21.7715 | -105.9556
21.7751 | -105.9491
21.7788 | -105.9411
21.7824 | -105.9345
21.786 | -105.9282
21.7896 | -105.9183
21.7932 | -105.9082
21.7968 | -105.8976
21.8004 | -105.8853
21.8041 | -105.8711
21.8077 | -105.859
21.8113 | -105.8475
21.8149 | -105.8356
21.8185 | -105.8237
21.8221 | -105.8116
21.8257 | -105.7998
21.8294 | -105.7877
21.833 | -105.775
21.8366 | -105.7626
21.8402 | -105.7513
21.8438 | -105.7396
21.8474 | -105.7275
21.851 | -105.7147
21.8546 | -105.702
21.8583 | -105.6939
21.8619 | -105.6853
21.8655 | -105.6761
21.8691 | -105.6685
21.8727 | -105.6611
21.8763 | -105.6518
21.8799 | -105.6431
21.8836 | -105.635
21.8872 | -105.6284
21.8908 | -105.6233
21.8944 | -105.6223
21.898 | -105.6162
21.9016 | -105.6112
21.9052 | -105.6077
21.9088 | -105.6064
21.9125 | -105.6029
21.9161 | -105.6001
21.9197 | -105.5966
21.9233 | -105.5933
21.9269 | -105.5919
21.9305 | -105.5925
21.9341 | -105.5911
21.9378 | -105.5902
21.9414 | -105.5897
21.945 | -105.5899
21.9486 | -105.5913
21.9522 | -105.5931
21.9558 | -105.5943
21.9594 | -105.5968
21.9631 | -105.5969
21.9667 | -105.5973
21.9703 | -105.5975
21.9739 | -105.5971
21.9775 | -105.598
21.9811 | -105.6
21.9847 | -105.6004
21.9883 | -105.603
21.992 | -105.6066
21.9956 | -105.6077
21.9992 | -105.6066
22.0028 | -105.6077
22.0064 | -105.6093
22.01 | -105.6121
22.0136 | -105.6144
22.0173 | -105.616
22.0209 | -105.6201
22.0245 | -105.6257
22.0281 | -105.6282
22.0317 | -105.6286
22.0353 | -105.6305
22.0389 | -105.6321
22.0425 | -105.6335
22.0462 | -105.6359
22.0498 | -105.638
22.0534 | -105.6384
22.057 | -105.6377
22.0606 | -105.6404
22.0642 | -105.6432
22.0678 | -105.6461
22.0715 | -105.6469
22.0751 | -105.6464
22.0787 | -105.6483
22.0823 | -105.6512
22.0859 | -105.6521
22.0895 | -105.6536
22.0931 | -105.6551
22.0968 | -105.6579
22.1004 | -105.6634
22.104 | -105.6679
22.1076 | -105.671
22.1112 | -105.6748
22.1148 | -105.681
22.1184 | -105.6841
22.122 | -105.6843
22.1257 | -105.6844
22.1293 | -105.6846
22.1329 | -105.6857
22.1365 | -105.6871
22.1401 | -105.6901
22.1437 | -105.6944
22.1473 | -105.6976
22.151 | -105.7003
22.1546 | -105.7044
22.1582 | -105.7074
22.1618 | -105.7116
22.1654 | -105.7158
22.169 | -105.7191
22.1726 | -105.723
22.1762 | -105.7271
22.1799 | -105.7313
22.1835 | -105.735
22.1871 | -105.7393
22.1907 | -105.7451
22.1943 | -105.7515
22.1979 | -105.7566
22.2015 | -105.7605
22.2052 | -105.7656
22.2088 | -105.7715
22.2124 | -105.7781
22.216 | -105.783
22.2196 | -105.7867
22.2232 | -105.7902
22.2268 | -105.7949
22.2305 | -105.7985
22.2341 | -105.8035
22.2377 | -105.8082
22.2413 | -105.8114
22.2449 | -105.8163
22.2485 | -105.8215
22.2521 | -105.8257
22.2557 | -105.8297
22.2594 | -105.8316
22.263 | -105.8327
22.2666 | -105.8346
22.2702 | -105.8364
22.2738 | -105.8393
22.2774 | -105.8437
22.281 | -105.8462
22.2847 | -105.848
22.2883 | -105.85
22.2919 | -105.8508
22.2955 | -105.8515
22.2991 | -105.8533
22.3027 | -105.8542
22.3063 | -105.852
22.31 | -105.8509
22.3136 | -105.8527
22.3172 | -105.8533
22.3208 | -105.8539
22.3244 | -105.8538
22.328 | -105.8517
22.3316 | -105.8506
22.3352 | -105.849
22.3389 | -105.8448
22.3425 | -105.8415
22.3461 | -105.84
22.3497 | -105.8404
22.3533 | -105.8414
22.3569 | -105.8404
22.3605 | -105.8378
22.3642 | -105.8361
22.3678 | -105.8352
22.3714 | -105.8316
22.375 | -105.827
22.3786 | -105.8232
22.3822 | -105.816
22.3858 | -105.8111
22.3894 | -105.8078
22.3931 | -105.8023
22.3967 | -105.7951
22.4003 | -105.7894
22.4039 | -105.7835
22.4075 | -105.7781
22.4111 | -105.7733
22.4147 | -105.7703
22.4184 | -105.7644
22.422 | -105.761
22.4256 | -105.758
22.4292 | -105.7533
22.4328 | -105.748
22.4364 | -105.7446
22.44 | -105.7452
22.4437 | -105.7464
22.4473 | -105.747
22.4509 | -105.7461
22.4545 | -105.7454
22.4581 | -105.7452
22.4617 | -105.7455
22.4653 | -105.7459
22.4689 | -105.7449
22.4726 | -105.7443
22.4762 | -105.7452
22.4798 | -105.7465
22.4834 | -105.7476
22.487 | -105.749
22.4906 | -105.7516
22.4942 | -105.7536
22.4979 | -105.7558
22.5015 | -105.7577
22.5051 | -105.7612
22.5087 | -105.7638
22.5123 | -105.7654
22.5159 | -105.7676
22.5195 | -105.7709
22.5231 | -105.7745
22.5268 | -105.7769
22.5304 | -105.78
22.534 | -105.7844
22.5376 | -105.7865
22.5412 | -105.7881
22.5448 | -105.7923
22.5484 | -105.7965
22.5521 | -105.7986
22.5557 | -105.799
22.5593 | -105.7994
22.5629 | -105.7986
22.5665 | -105.7992
22.5701 | -105.8
22.5737 | -105.7998
22.5774 | -105.8
22.581 | -105.8003
22.5846 | -105.7993
22.5882 | -105.797
22.5918 | -105.7935
22.5954 | -105.7888
22.599 | -105.7858
22.6026 | -105.7831
22.6063 | -105.7801
22.6099 | -105.7773
22.6135 | -105.7728
22.6171 | -105.7679
22.6207 | -105.7631
22.6243 | -105.7598
22.6279 | -105.7569
22.6316 | -105.7523
22.6352 | -105.7468
22.6388 | -105.7419
22.6424 | -105.7372
22.646 | -105.7347
22.6496 | -105.7313
22.6532 | -105.7258
22.6568 | -105.7218
22.6605 | -105.7172
22.6641 | -105.7146
22.6677 | -105.7113
22.6713 | -105.707
22.6749 | -105.7044
22.6785 | -105.7019
22.6821 | -105.6987
22.6858 | -105.697
22.6894 | -105.6988
22.693 | -105.6984
22.6966 | -105.6977
22.7002 | -105.7007
22.7038 | -105.7024
22.7074 | -105.7034
22.7111 | -105.704
22.7147 | -105.7059
22.7183 | -105.7082
22.7219 | -105.7115
22.7255 | -105.7151
22.7291 | -105.7165
22.7327 | -105.7186
22.7363 | -105.7221
22.74 | -105.7242
22.7436 | -105.7248
22.7472 | -105.7262
22.7508 | -105.7299
22.7544 | -105.7351
22.758 | -105.738
22.7616 | -105.7387
22.7653 | -105.7391
22.7689 | -105.7398
22.7725 | -105.7429
22.7761 | -105.7462
22.7797 | -105.7471
22.7833 | -105.7466
22.7869 | -105.7438
22.7906 | -105.7419
22.7942 | -105.7411
22.7978 | -105.7388
22.8014 | -105.7365
22.805 | -105.7357
22.8086 | -105.7311
22.8122 | -105.7246
22.8158 | -105.7202
22.8195 | -105.714
22.8231 | -105.7072
22.8267 | -105.6988
22.8303 | -105.6898
22.8339 | -105.6804
22.8375 | -105.6708
22.8411 | -105.661
22.8448 | -105.6512
22.8484 | -105.6414
22.852 | -105.63
22.8556 | -105.6185
22.8592 | -105.6061
22.8628 | -105.5945
22.8664 | -105.5815
22.87 | -105.568
22.8737 | -105.5545
22.8773 | -105.5427
22.8809 | -105.531
22.8845 | -105.5186
22.8881 | -105.5055
22.8917 | -105.4903
22.8953 | -105.4766
22.899 | -105.4643
22.9026 | -105.4524
22.9062 | -105.4395
22.9098 | -105.428
22.9134 | -105.4135
22.917 | -105.3999
22.9206 | -105.3874
22.9243 | -105.3765
22.9279 | -105.3665
22.9315 | -105.3554
22.9351 | -105.3454
22.9387 | -105.3362
22.9423 | -105.3272
22.9459 | -105.3171
22.9495 | -105.3085
22.9532 | -105.2983
22.9568 | -105.2877
22.9604 | -105.2763
22.964 | -105.2673
22.9676 | -105.2591
22.9712 | -105.2509
22.9748 | -105.2422
22.9785 | -105.2361
22.9821 | -105.2316
22.9857 | -105.2263
22.9893 | -105.219
22.9929 | -105.2128
22.9965 | -105.2091
23.0001 | -105.2046
23.0037 | -105.2009
23.0074 | -105.1979
23.011 | -105.1916
23.0146 | -105.1863
23.0182 | -105.1844
23.0218 | -105.1831
23.0254 | -105.1803
23.029 | -105.1779
23.0327 | -105.1768
23.0363 | -105.1766
23.0399 | -105.1767
23.0435 | -105.1776
23.0471 | -105.1795
23.0507 | -105.1813
23.0543 | -105.1838
23.058 | -105.1874
23.0616 | -105.1913
23.0652 | -105.1972
23.0688 | -105.2041
23.0724 | -105.2112
23.076 | -105.2181
23.0796 | -105.2236
23.0832 | -105.2315
23.0869 | -105.2396
23.0905 | -105.2471
23.0941 | -105.256
23.0977 | -105.2646
23.1013 | -105.2722
23.1049 | -105.2811
23.1085 | -105.2919
23.1122 | -105.3038
23.1158 | -105.314
23.1194 | -105.3228
23.123 | -105.3322
23.1266 | -105.3407
23.1302 | -105.348
23.1338 | -105.356
23.1375 | -105.3653
23.1411 | -105.3738
23.1447 | -105.3825
23.1483 | -105.3914
23.1519 | -105.3995
23.1555 | -105.4071
23.1591 | -105.4142
23.1627 | -105.4188
23.1664 | -105.422
23.17 | -105.4245
23.1736 | -105.4274
23.1772 | -105.4306
23.1808 | -105.433
23.1844 | -105.4371
23.188 | -105.4402
23.1917 | -105.4431
23.1953 | -105.4456
23.1989 | -105.4488
23.2025 | -105.451
23.2061 | -105.4528
23.2097 | -105.4548
23.2133 | -105.4581
23.2169 | -105.4613
23.2206 | -105.4633
23.2242 | -105.466
23.2278 | -105.4664
23.2314 | -105.4664
23.235 | -105.4686
23.2386 | -105.4733
23.2422 | -105.4764
23.2459 | -105.4779
23.2495 | -105.479
23.2531 | -105.4806
23.2567 | -105.4825
23.2603 | -105.487
23.2639 | -105.4888
23.2675 | -105.4925
23.2711 | -105.4972
23.2748 | -105.5029
23.2784 | -105.5046
23.282 | -105.5076
23.2856 | -105.5111
23.2892 | -105.5159
23.2928 | -105.5208
23.2964 | -105.5253
23.3001 | -105.5278
23.3037 | -105.5324
23.3073 | -105.5353
23.3109 | -105.5371
23.3145 | -105.5417
23.3181 | -105.5468
23.3217 | -105.5507
23.3254 | -105.5532
23.329 | -105.5559
23.3326 | -105.5597
23.3362 | -105.5634
23.3398 | -105.5705
23.3434 | -105.576
23.347 | -105.5782
23.3506 | -105.5795
23.3543 | -105.5818
23.3579 | -105.585
23.3615 | -105.5862
23.3651 | -105.5863
23.3687 | -105.5856
23.3723 | -105.5831
23.3759 | -105.5788
23.3796 | -105.5759
23.3832 | -105.5744
23.3868 | -105.572
23.3904 | -105.5688
23.394 | -105.5668
23.3976 | -105.5645
23.4012 | -105.5618
23.4049 | -105.5594
23.4085 | -105.5575
23.4121 | -105.5546
23.4157 | -105.5521
23.4193 | -105.5509
23.4229 | -105.5492
23.4265 | -105.5481
23.4301 | -105.5475
23.4338 | -105.5475
23.4374 | -105.5473
23.441 | -105.5488
23.4446 | -105.5522
23.4482 | -105.5538
23.4518 | -105.5546
23.4554 | -105.5555
23.4591 | -105.5625
23.4627 | -105.5652
23.4663 | -105.5676
23.4699 | -105.5715
23.4735 | -105.5768
23.4771 | -105.5808
23.4807 | -105.5834
23.4843 | -105.5897
23.488 | -105.5969
23.4916 | -105.6034
23.4952 | -105.6103
23.4988 | -105.618
23.5024 | -105.6263
23.506 | -105.6333
23.5096 | -105.6418
23.5133 | -105.651
23.5169 | -105.6604
23.5205 | -105.6696
23.5241 | -105.6779
23.5277 | -105.6862
23.5313 | -105.6949
23.5349 | -105.7031
23.5386 | -105.7125
23.5422 | -105.7219
23.5458 | -105.73
23.5494 | -105.7416
23.553 | -105.7528
23.5566 | -105.7613
23.5602 | -105.7744
23.5638 | -105.7882
23.5675 | -105.7984
23.5711 | -105.8076
23.5747 | -105.8177
23.5783 | -105.8284
23.5819 | -105.8378
23.5855 | -105.8466
23.5891 | -105.8562
23.5928 | -105.8654
23.5964 | -105.8749
23.6 | -105.8827
23.6036 | -105.8892
23.6072 | -105.8967
23.6108 | -105.9074
23.6144 | -105.9168
23.6181 | -105.9263
23.6217 | -105.9354
23.6253 | -105.9424
23.6289 | -105.9507
23.6325 | -105.9602
23.6361 | -105.9695
23.6397 | -105.9792
23.6433 | -105.9876
23.647 | -105.994
23.6506 | -106.0023
23.6542 | -106.0123
23.6578 | -106.0219
23.6614 | -106.0315
23.665 | -106.0408
23.6686 | -106.0501
23.6723 | -106.0596
23.6759 | -106.0697
23.6795 | -106.0806
23.6831 | -106.0927
23.6867 | -106.1058
23.6903 | -106.1181
23.6939 | -106.1326
23.6975 | -106.1481
23.7012 | -106.1649
23.7048 | -106.1818
23.7084 | -106.1997
23.712 | -106.2167
23.7156 | -106.2331
23.7192 | -106.2492
23.7228 | -106.2655
23.7265 | -106.2802
23.7301 | -106.2958
23.7337 | -106.3132
23.7373 | -106.3305
23.7409 | -106.3454
23.7445 | -106.3621
23.7481 | -106.3799
23.7518 | -106.3961
23.7554 | -106.413
23.759 | -106.429
23.7626 | -106.4455
23.7662 | -106.46
23.7698 | -106.4738
23.7734 | -106.4873
23.777 | -106.4985
23.7807 | -106.5124
23.7843 | -106.5258
23.7879 | -106.5381
23.7915 | -106.5497
23.7951 | -106.5576
23.7987 | -106.5653
23.8023 | -106.575
23.806 | -106.5823
23.8096 | -106.5872
23.8132 | -106.5934
23.8168 | -106.6007
23.8204 | -106.6031
23.824 | -106.6044
23.8276 | -106.6072
23.8312 | -106.611
23.8349 | -106.6133
23.8385 | -106.6132
23.8421 | -106.6128
23.8457 | -106.6116
23.8493 | -106.6086
23.8529 | -106.6037
23.8565 | -106.5991
23.8602 | -106.5952
23.8638 | -106.5903
23.8674 | -106.5845
23.871 | -106.5791
23.8746 | -106.5736
23.8782 | -106.5673
23.8818 | -106.5625
23.8855 | -106.5541
23.8891 | -106.5457
23.8927 | -106.539
23.8963 | -106.5269
23.8999 | -106.5155
23.9035 | -106.5069
23.9071 | -106.4992
23.9107 | -106.4912
23.9144 | -106.4815
23.918 | -106.4718
23.9216 | -106.4623
23.9252 | -106.4543
23.9288 | -106.4447
23.9324 | -106.434
23.936 | -106.4232
23.9397 | -106.4123
23.9433 | -106.4023
23.9469 | -106.3923
23.9505 | -106.3834
23.9541 | -106.373
23.9577 | -106.3606
23.9613 | -106.3502
23.9649 | -106.3406
23.9686 | -106.3304
23.9722 | -106.3198
23.9758 | -106.3085
23.9794 | -106.3001
23.983 | -106.2905
23.9866 | -106.2817
23.9902 | -106.2712
23.9939 | -106.261
23.9975 | -106.2525
24.0011 | -106.2456
24.0047 | -106.2396
24.0083 | -106.234
24.0119 | -106.2278
24.0155 | -106.2214
24.0192 | -106.2169
24.0228 | -106.2136
24.0264 | -106.2094
24.03 | -106.2072
24.0336 | -106.2037
24.0372 | -106.2018
24.0408 | -106.2006
24.0444 | -106.2017
24.0481 | -106.2031
24.0517 | -106.2041
24.0553 | -106.2045
24.0589 | -106.2068
24.0625 | -106.2088
24.0661 | -106.2094
24.0697 | -106.2115
24.0734 | -106.2146
24.077 | -106.2181
24.0806 | -106.2208
24.0842 | -106.2228
24.0878 | -106.2262
24.0914 | -106.2313
24.095 | -106.2357
24.0986 | -106.2421
24.1023 | -106.2491
24.1059 | -106.2554
24.1095 | -106.2617
24.1131 | -106.2674
24.1167 | -106.2733
24.1203 | -106.2791
24.1239 | -106.2858
24.1276 | -106.2907
24.1312 | -106.2934
24.1348 | -106.2971
24.1384 | -106.302
24.142 | -106.3062
24.1456 | -106.3072
24.1492 | -106.3074
24.1529 | -106.3072
24.1565 | -106.3078
24.1601 | -106.3077
24.1637 | -106.3064
24.1673 | -106.3042
24.1709 | -106.3019
24.1745 | -106.2994
24.1781 | -106.296
24.1818 | -106.2917
24.1854 | -106.2853
24.189 | -106.2791
24.1926 | -106.2729
24.1962 | -106.2659
24.1998 | -106.2589
24.2034 | -106.2514
24.2071 | -106.2437
24.2107 | -106.2351
24.2143 | -106.224
24.2179 | -106.2145
24.2215 | -106.2042
24.2251 | -106.1936
24.2287 | -106.1837
24.2324 | -106.1725
24.236 | -106.1621
24.2396 | -106.1497
24.2432 | -106.1338
24.2468 | -106.1207
24.2504 | -106.1062
24.254 | -106.0887
24.2576 | -106.0672
24.2613 | -106.0312
24.2649 | -105.9859
24.2685 | -105.944
24.2721 | -105.9049
24.2757 | -105.8652
24.2793 | -105.826
24.2829 | -105.7882
24.2866 | -105.7484
24.2902 | -105.705
24.2938 | -105.6582
24.2974 | -105.6127
24.301 | -105.5608
24.3046 | -105.5085
24.3082 | -105.4574
24.3118 | -105.407
24.3155 | -105.3568
24.3191 | -105.3037
24.3227 | -105.2493
24.3263 | -105.1936
24.3299 | -105.1351
24.3335 | -105.0762
24.3371 | -105.0154
24.3408 | -104.9533
24.3444 | -104.8912
24.348 | -104.8278
24.3516 | -104.7635
24.3552 | -104.6985
24.3588 | -104.634
24.3624 | -104.5688
24.3661 | -104.503
24.3697 | -104.4393
24.3733 | -104.3778
24.3769 | -104.3175
24.3805 | -104.2567
24.3841 | -104.1981
24.3877 | -104.141
24.3913 | -104.0868
24.395 | -104.036
24.3986 | -103.9889
24.4022 | -103.9442
24.4058 | -103.9032
24.4094 | -103.8663
24.413 | -103.8313
24.4166 | -103.7985
24.4203 | -103.7689
24.4239 | -103.7444
24.4275 | -103.7241
24.4311 | -103.7067
24.4347 | -103.6933
24.4383 | -103.6831
24.4419 | -103.6763
24.4455 | -103.6748
24.4492 | -103.6782
24.4528 | -103.6855
24.4564 | -103.697
24.46 | -103.7131
24.4636 | -103.732
24.4672 | -103.7548
24.4708 | -103.7811
24.4745 | -103.8128
24.4781 | -103.8472
24.4817 | -103.8838
24.4853 | -103.9246
24.4889 | -103.9686
24.4925 | -104.0143
24.4961 | -104.0633
24.4998 | -104.1273
24.5034 | -104.1774
24.507 | -104.2289
24.5106 | -104.2827
24.5142 | -104.3359
24.5178 | -104.3893
24.5214 | -104.4431
24.525 | -104.4976
24.5287 | -104.553
24.5323 | -104.6067
24.5359 | -104.6588
24.5395 | -104.7071
24.5431 | -104.7525
24.5467 | -104.7946
24.5503 | -104.8367
24.554 | -104.876
24.5576 | -104.9095
24.5612 | -104.9447
24.5648 | -104.9772
24.5684 | -105.0074
24.572 | -105.0351
24.5756 | -105.0562
24.5792 | -105.073
24.5829 | -105.0864
24.5865 | -105.0984
24.5901 | -105.1089
24.5937 | -105.1137
24.5973 | -105.1167
24.6009 | -105.1186
24.6045 | -105.1184
24.6082 | -105.1223
24.6118 | -105.1255
24.6154 | -105.1206
24.619 | -105.1077
24.6226 | -105.0913
24.6262 | -105.0708
24.6298 | -105.0447
24.6335 | -105.0149
24.6371 | -104.9841
24.6407 | -104.9539
24.6443 | -104.9229
24.6479 | -104.8931
24.6515 | -104.8596
24.6551 | -104.8247
24.6587 | -104.7878
24.6624 | -104.7517
24.666 | -104.7171
24.6696 | -104.6839
24.6732 | -104.6519
24.6768 | -104.6225
24.6804 | -104.5944
24.684 | -104.5676
24.6877 | -104.5426
24.6913 | -104.5195
24.6949 | -104.4996
24.6985 | -104.4825
24.7021 | -104.4688
24.7057 | -104.4575
24.7093 | -104.4487
24.7129 | -104.4431
24.7166 | -104.4393
24.7202 | -104.4372
24.7238 | -104.438
24.7274 | -104.4424
24.731 | -104.4476
24.7346 | -104.4543
24.7382 | -104.4645
24.7419 | -104.4777
24.7455 | -104.4929
24.7491 | -104.5097
24.7527 | -104.5292
24.7563 | -104.5519
24.7599 | -104.578
24.7635 | -104.6069
24.7672 | -104.6379
24.7708 | -104.6705
24.7744 | -104.7029
24.778 | -104.7371
24.7816 | -104.7727
24.7852 | -104.8102
24.7888 | -104.8479
24.7924 | -104.8859
24.7961 | -104.925
24.7997 | -104.9649
24.8033 | -105.0047
24.8069 | -105.0441
24.8105 | -105.0847
24.8141 | -105.1254
24.8177 | -105.1668
24.8214 | -105.2051
24.825 | -105.2429
24.8286 | -105.2811
24.8322 | -105.3167
24.8358 | -105.3529
24.8394 | -105.3892
24.843 | -105.4212
24.8467 | -105.4523
24.8503 | -105.4818
24.8539 | -105.5117
24.8575 | -105.5413
24.8611 | -105.5679
24.8647 | -105.5949
24.8683 | -105.6199
24.8719 | -105.6434
24.8756 | -105.6656
24.8792 | -105.687
24.8828 | -105.7075
24.8864 | -105.7253
24.89 | -105.7456
24.8936 | -105.7645
24.8972 | -105.7824
24.9009 | -105.7973
24.9045 | -105.808
24.9081 | -105.8219
24.9117 | -105.8308
24.9153 | -105.8362
24.9189 | -105.8389
24.9225 | -105.8412
24.9261 | -105.8464
24.9298 | -105.8524
24.9334 | -105.8572
24.937 | -105.86
24.9406 | -105.8603
24.9442 | -105.8596
24.9478 | -105.8587
24.9514 | -105.8573
24.9551 | -105.856
24.9587 | -105.8556
24.9623 | -105.8551
24.9659 | -105.855
24.9695 | -105.8525
24.9731 | -105.8479
24.9767 | -105.8442
24.9804 | -105.8397
24.984 | -105.8356
24.9876 | -105.8318
24.9912 | -105.8274
24.9948 | -105.822
24.9984 | -105.8162
25.002 | -105.8118
25.0056 | -105.8065
25.0093 | -105.801
25.0129 | -105.7941
25.0165 | -105.787
25.0201 | -105.778
25.0237 | -105.7695
25.0273 | -105.7627
25.0309 | -105.7568
25.0346 | -105.7501
25.0382 | -105.744
25.0418 | -105.7383
25.0454 | -105.7305
25.049 | -105.722
25.0526 | -105.7159
25.0562 | -105.7091
25.0598 | -105.703
25.0635 | -105.6974
25.0671 | -105.6909
25.0707 | -105.6843
25.0743 | -105.6784
25.0779 | -105.6743
25.0815 | -105.6695
25.0851 | -105.6656
25.0888 | -105.663
25.0924 | -105.6601
25.096 | -105.6554
25.0996 | -105.6503
25.1032 | -105.6483
25.1068 | -105.6475
25.1104 | -105.6446
25.1141 | -105.6419
25.1177 | -105.6399
25.1213 | -105.6391
25.1249 | -105.6393
25.1285 | -105.6407
25.1321 | -105.6417
25.1357 | -105.6428
25.1393 | -105.6449
25.143 | -105.6499
25.1466 | -105.6528
25.1502 | -105.6544
25.1538 | -105.6585
25.1574 | -105.6639
25.161 | -105.6697
25.1646 | -105.6748
25.1683 | -105.679
25.1719 | -105.6832
25.1755 | -105.6902
25.1791 | -105.6981
25.1827 | -105.705
25.1863 | -105.7118
25.1899 | -105.7183
25.1935 | -105.7254
25.1972 | -105.7336
25.2008 | -105.7406
25.2044 | -105.7485
25.208 | -105.7558
25.2116 | -105.7634
25.2152 | -105.7718
25.2188 | -105.7785
25.2225 | -105.7851
25.2261 | -105.7913
25.2297 | -105.7964
25.2333 | -105.7998
25.2369 | -105.8058
25.2405 | -105.8129
25.2441 | -105.8184
25.2478 | -105.8229
25.2514 | -105.8281
25.255 | -105.8333
25.2586 | -105.8382
25.2622 | -105.8412
25.2658 | -105.8427
25.2694 | -105.8449
25.273 | -105.8488
25.2767 | -105.8527
25.2803 | -105.8556
25.2839 | -105.8579
25.2875 | -105.86
25.2911 | -105.8626
25.2947 | -105.8638
25.2983 | -105.8625
25.302 | -105.8603
25.3056 | -105.859
25.3092 | -105.8587
25.3128 | -105.8578
25.3164 | -105.8567
25.32 | -105.8538
25.3236 | -105.8514
25.3273 | -105.8474
25.3309 | -105.8424
25.3345 | -105.8355
25.3381 | -105.83
25.3417 | -105.8259
25.3453 | -105.8211
25.3489 | -105.8134
25.3525 | -105.8049
25.3562 | -105.795
25.3598 | -105.787
25.3634 | -105.78
25.367 | -105.7718
25.3706 | -105.7641
25.3742 | -105.7559
25.3778 | -105.7478
25.3815 | -105.7401
25.3851 | -105.7302
25.3887 | -105.7185
25.3923 | -105.7076
25.3959 | -105.6973
25.3995 | -105.6876
25.4031 | -105.6781
25.4067 | -105.6688
25.4104 | -105.6584
25.414 | -105.6482
25.4176 | -105.6376
25.4212 | -105.6282
25.4248 | -105.6179
25.4284 | -105.6087
25.432 | -105.5999
25.4357 | -105.5889
25.4393 | -105.5801
25.4429 | -105.5738
25.4465 | -105.5658
25.4501 | -105.5561
25.4537 | -105.5493
25.4573 | -105.544
25.461 | -105.5388
25.4646 | -105.5324
25.4682 | -105.5279
25.4718 | -105.5231
25.4754 | -105.519
25.479 | -105.5148
25.4826 | -105.5113
25.4862 | -105.5094
25.4899 | -105.5075
25.4935 | -105.5063
25.4971 | -105.5054
25.5007 | -105.5031
25.5043 | -105.5017
25.5079 | -105.5017
25.5115 | -105.5017
25.5152 | -105.5012
25.5188 | -105.503
25.5224 | -105.5048
25.526 | -105.5072
25.5296 | -105.5108
25.5332 | -105.5143
25.5368 | -105.5173
25.5404 | -105.5207
25.5441 | -105.5259
25.5477 | -105.5293
25.5513 | -105.5331
25.5549 | -105.5355
25.5585 | -105.538
25.5621 | -105.5421
25.5657 | -105.5454
25.5694 | -105.5473
25.573 | -105.5505
25.5766 | -105.5533
25.5802 | -105.5559
25.5838 | -105.5592
25.5874 | -105.5641
25.591 | -105.5692
25.5947 | -105.5735
25.5983 | -105.5759
25.6019 | -105.5791
25.6055 | -105.5819
25.6091 | -105.5847
25.6127 | -105.5875
25.6163 | -105.5882
25.6199 | -105.5909
25.6236 | -105.5944
25.6272 | -105.5971
25.6308 | -105.6001
25.6344 | -105.602
25.638 | -105.6025
25.6416 | -105.6044
25.6452 | -105.6052
25.6489 | -105.6052
25.6525 | -105.6044
25.6561 | -105.606
25.6597 | -105.6096
25.6633 | -105.6115
25.6669 | -105.6132
25.6705 | -105.6127
25.6742 | -105.6122
25.6778 | -105.6124
25.6814 | -105.6129
25.685 | -105.614
25.6886 | -105.6182
25.6922 | -105.6213
25.6958 | -105.6249
25.6994 | -105.6276
25.7031 | -105.6286
25.7067 | -105.6296
25.7103 | -105.6316
25.7139 | -105.6352
25.7175 | -105.6385
25.7211 | -105.6391
25.7247 | -105.6401
25.7284 | -105.6428
25.732 | -105.645
25.7356 | -105.6485
25.7392 | -105.6518
25.7428 | -105.656
25.7464 | -105.6608
25.75 | -105.6657
25.7536 | -105.6705
25.7573 | -105.6746
25.7609 | -105.6792
25.7645 | -105.6849
25.7681 | -105.6884
25.7717 | -105.6925
25.7753 | -105.6976
25.7789 | -105.7028
25.7826 | -105.707
25.7862 | -105.7129
25.7898 | -105.7179
25.7934 | -105.7243
25.797 | -105.733
25.8006 | -105.7416
25.8042 | -105.7473
25.8078 | -105.7541
25.8115 | -105.7615
25.8151 | -105.7673
25.8187 | -105.773
25.8223 | -105.7765
25.8259 | -105.7793
25.8295 | -105.7843
25.8331 | -105.7897
25.8368 | -105.7953
25.8404 | -105.8018
25.844 | -105.8067
25.8476 | -105.8106
25.8512 | -105.8161
25.8548 | -105.8219
25.8584 | -105.8277
25.8621 | -105.8344
25.8657 | -105.842
25.8693 | -105.8485
25.8729 | -105.8538
25.8765 | -105.86
25.8801 | -105.8647
25.8837 | -105.8697
25.8873 | -105.8755
25.891 | -105.8793
25.8946 | -105.8825
25.8982 | -105.8876
25.9018 | -105.8922
25.9054 | -105.8966
25.909 | -105.9022
25.9126 | -105.9079
25.9163 | -105.9127
25.9199 | -105.9191
25.9235 | -105.9255
25.9271 | -105.9299
25.9307 | -105.9349
25.9343 | -105.9402
25.9379 | -105.9455
25.9416 | -105.951
25.9452 | -105.9575
25.9488 | -105.9638
25.9524 | -105.9686
25.956 | -105.9749
25.9596 | -105.981
25.9632 | -105.988
25.9668 | -105.995
25.9705 | -106.0017
25.9741 | -106.0083
25.9777 | -106.0146
25.9813 | -106.0234
25.9849 | -106.0317
25.9885 | -106.039
25.9921 | -106.0463
25.9958 | -106.0524
25.9994 | -106.059
26.003 | -106.0692
26.0066 | -106.0783
26.0102 | -106.0879
26.0138 | -106.0974
26.0174 | -106.1058
26.021 | -106.1134
26.0247 | -106.1212
26.0283 | -106.1316
26.0319 | -106.1414
26.0355 | -106.1528
26.0391 | -106.1639
26.0427 | -106.1748
26.0463 | -106.1848
26.05 | -106.1952
26.0536 | -106.2026
26.0572 | -106.2102
26.0608 | -106.2184
26.0644 | -106.2282
26.068 | -106.2378
26.0716 | -106.2484
26.0753 | -106.2574
26.0789 | -106.2663
26.0825 | -106.2736
26.0861 | -106.2832
26.0897 | -106.2936
26.0933 | -106.3033
26.0969 | -106.3117
26.1005 | -106.3197
26.1042 | -106.327
26.1078 | -106.3338
26.1114 | -106.3392
26.115 | -106.3449
26.1186 | -106.3517
26.1222 | -106.3587
26.1258 | -106.3657
26.1295 | -106.3722
26.1331 | -106.3774
26.1367 | -106.3848
26.1403 | -106.3904
26.1439 | -106.3941
26.1475 | -106.3992
26.1511 | -106.4047
26.1547 | -106.4086
26.1584 | -106.4128
26.162 | -106.4143
26.1656 | -106.4153
26.1692 | -106.416
26.1728 | -106.4168
26.1764 | -106.4179
26.18 | -106.4199
26.1837 | -106.4209
26.1873 | -106.4208
26.1909 | -106.4213
26.1945 | -106.4231
26.1981 | -106.4243
26.2017 | -106.4245
26.2053 | -106.4242
26.209 | -106.4226
26.2126 | -106.4224
26.2162 | -106.4222
26.2198 | -106.4217
26.2234 | -106.419
26.227 | -106.4163
26.2306 | -106.4132
26.2342 | -106.4102
26.2379 | -106.4084
26.2415 | -106.4037
26.2451 | -106.3985
26.2487 | -106.3943
26.2523 | -106.3911
26.2559 | -106.3895
26.2595 | -106.3885
26.2632 | -106.3879
26.2668 | -106.386
26.2704 | -106.3848
26.274 | -106.3842
26.2776 | -106.382
26.2812 | -106.3807
26.2848 | -106.3801
26.2885 | -106.3789
26.2921 | -106.3781
26.2957 | -106.3786
26.2993 | -106.38
26.3029 | -106.3817
26.3065 | -106.3827
26.3101 | -106.3836
26.3137 | -106.3851
26.3174 | -106.3863
26.321 | -106.3877
26.3246 | -106.389
26.3282 | -106.3908
26.3318 | -106.3937
26.3354 | -106.3967
26.339 | -106.4007
26.3427 | -106.4072
26.3463 | -106.4125
26.3499 | -106.4173
26.3535 | -106.421
26.3571 | -106.4252
26.3607 | -106.4318
26.3643 | -106.4393
26.3679 | -106.4461
26.3716 | -106.4503
26.3752 | -106.4554
26.3788 | -106.4616
26.3824 | -106.4705
26.386 | -106.4788
26.3896 | -106.4856
26.3932 | -106.4923
26.3969 | -106.4985
26.4005 | -106.5046
26.4041 | -106.5112
26.4077 | -106.5175
26.4113 | -106.5209
26.4149 | -106.5238
26.4185 | -106.5289
26.4222 | -106.535
26.4258 | -106.5412
26.4294 | -106.5483
26.433 | -106.554
26.4366 | -106.5582
26.4402 | -106.5607
26.4438 | -106.5639
26.4474 | -106.568
26.4511 | -106.5715
26.4547 | -106.576
26.4583 | -106.5792
26.4619 | -106.5823
26.4655 | -106.5838
26.4691 | -106.5868
26.4727 | -106.5894
26.4764 | -106.5907
26.48 | -106.5919
26.4836 | -106.593
26.4872 | -106.5929
26.4908 | -106.5912
26.4944 | -106.5894
26.498 | -106.5876
26.5016 | -106.5848
26.5053 | -106.5834
26.5089 | -106.5818
26.5125 | -106.581
26.5161 | -106.5801
26.5197 | -106.5763
26.5233 | -106.5725
26.5269 | -106.5705
26.5306 | -106.5674
26.5342 | -106.565
26.5378 | -106.5618
26.5414 | -106.5566
26.545 | -106.5511
26.5486 | -106.5464
26.5522 | -106.5413
26.5559 | -106.5355
26.5595 | -106.5306
26.5631 | -106.5262
26.5667 | -106.5211
26.5703 | -106.5131
26.5739 | -106.5051
26.5775 | -106.4987
26.5811 | -106.493
26.5848 | -106.4843
26.5884 | -106.4764
26.592 | -106.4687
26.5956 | -106.4606
26.5992 | -106.4539
26.6028 | -106.4458
26.6064 | -106.4389
26.6101 | -106.4323
26.6137 | -106.4261
26.6173 | -106.422
26.6209 | -106.4172
26.6245 | -106.4112
26.6281 | -106.4045
26.6317 | -106.3964
26.6353 | -106.3897
26.639 | -106.3826
26.6426 | -106.3769
26.6462 | -106.372
26.6498 | -106.3662
26.6534 | -106.3609
26.657 | -106.3571
26.6606 | -106.3535
26.6643 | -106.3501
26.6679 | -106.3459
26.6715 | -106.3418
26.6751 | -106.3384
26.6787 | -106.3336
26.6823 | -106.3289
26.6859 | -106.3272
26.6896 | -106.3245
26.6932 | -106.3211
26.6968 | -106.3183
26.7004 | -106.3148
26.704 | -106.3137
26.7076 | -106.3131
26.7112 | -106.3128
26.7148 | -106.3132
26.7185 | -106.3117
26.7221 | -106.3111
26.7257 | -106.3105
26.7293 | -106.3129
26.7329 | -106.3155
26.7365 | -106.3164
26.7401 | -106.3174
26.7438 | -106.3191
26.7474 | -106.3216
26.751 | -106.3236
26.7546 | -106.3245
26.7582 | -106.3262
26.7618 | -106.3289
26.7654 | -106.3334
26.7691 | -106.337
26.7727 | -106.3398
26.7763 | -106.3435
26.7799 | -106.3483
26.7835 | -106.3525
26.7871 | -106.3586
26.7907 | -106.3657
26.7943 | -106.3707
26.798 | -106.3747
26.8016 | -106.3795
26.8052 | -106.3845
26.8088 | -106.3892
26.8124 | -106.3958
26.816 | -106.4049
26.8196 | -106.4125
26.8233 | -106.4199
26.8269 | -106.4274
26.8305 | -106.4346
26.8341 | -106.441
26.8377 | -106.4474
26.8413 | -106.4542
26.8449 | -106.4607
26.8485 | -106.4676
26.8522 | -106.4752
26.8558 | -106.4833
26.8594 | -106.4917
26.863 | -106.5
26.8666 | -106.5085
26.8702 | -106.5152
26.8738 | -106.5216
26.8775 | -106.5286
26.8811 | -106.5347
26.8847 | -106.5419
26.8883 | -106.5482
26.8919 | -106.5539
26.8955 | -106.5587
26.8991 | -106.5626
26.9028 | -106.5661
26.9064 | -106.5698
26.91 | -106.5739
26.9136 | -106.5782
26.9172 | -106.5834
26.9208 | -106.5898
26.9244 | -106.5958
26.928 | -106.599
26.9317 | -106.6021
26.9353 | -106.6054
26.9389 | -106.6085
26.9425 | -106.612
26.9461 | -106.6163
26.9497 | -106.6212
26.9533 | -106.623
26.957 | -106.6238
26.9606 | -106.6255
26.9642 | -106.6265
26.9678 | -106.6281
26.9714 | -106.6296
26.975 | -106.6305
26.9786 | -106.631
26.9822 | -106.6321
26.9859 | -106.6314
26.9895 | -106.6295
26.9931 | -106.6281
26.9967 | -106.6278
27.0003 | -106.6279
27.0039 | -106.6281
27.0075 | -106.626
27.0112 | -106.6235
27.0148 | -106.6203
27.0184 | -106.6182
27.022 | -106.6153
27.0256 | -106.6121
27.0292 | -106.6088
27.0328 | -106.6054
27.0365 | -106.6008
27.0401 | -106.597
27.0437 | -106.594
27.0473 | -106.5904
27.0509 | -106.5879
27.0545 | -106.5844
27.0581 | -106.5827
27.0617 | -106.5805
27.0654 | -106.5785
27.069 | -106.5777
27.0726 | -106.578
27.0762 | -106.5788
27.0798 | -106.5782
27.0834 | -106.5764
27.087 | -106.5747
27.0907 | -106.5722
27.0943 | -106.5698
27.0979 | -106.5678
27.1015 | -106.5673
27.1051 | -106.5724
27.1087 | -106.5738
27.1123 | -106.5734
27.1159 | -106.573
27.1196 | -106.5732
27.1232 | -106.5755
27.1268 | -106.5774
27.1304 | -106.5786
27.134 | -106.5805
27.1376 | -106.5837
27.1412 | -106.586
27.1449 | -106.5896
27.1485 | -106.5923
27.1521 | -106.5968
27.1557 | -106.6009
27.1593 | -106.6048
27.1629 | -106.6089
27.1665 | -106.6146
27.1702 | -106.6202
27.1738 | -106.6254
27.1774 | -106.6304
27.181 | -106.6379
27.1846 | -106.6446
27.1882 | -106.6499
27.1918 | -106.6556
27.1954 | -106.6627
27.1991 | -106.6676
27.2027 | -106.6733
27.2063 | -106.68
27.2099 | -106.6865
27.2135 | -106.694
27.2171 | -106.7016
27.2207 | -106.709
27.2244 | -106.7165
27.228 | -106.7232
27.2316 | -106.7301
27.2352 | -106.7368
27.2388 | -106.7428
27.2424 | -106.7498
27.246 | -106.757
27.2496 | -106.7636
27.2533 | -106.7713
27.2569 | -106.7782
27.2605 | -106.7846
27.2641 | -106.7908
27.2677 | -106.7954
27.2713 | -106.801
27.2749 | -106.8071
27.2786 | -106.8142
27.2822 | -106.8208
27.2858 | -106.8252
27.2894 | -106.8295
27.293 | -106.8339
27.2966 | -106.8379
27.3002 | -106.8408
27.3039 | -106.8434
27.3075 | -106.8453
27.3111 | -106.8477
27.3147 | -106.85
27.3183 | -106.8507
27.3219 | -106.8511
27.3255 | -106.8511
27.3291 | -106.8503
27.3328 | -106.8491
27.3364 | -106.849
27.34 | -106.849
27.3436 | -106.8487
27.3472 | -106.8503
27.3508 | -106.8496
27.3544 | -106.8479
27.3581 | -106.8467
27.3617 | -106.8448
27.3653 | -106.8416
27.3689 | -106.8378
27.3725 | -106.834
27.3761 | -106.8295
27.3797 | -106.8228
27.3834 | -106.8153
27.387 | -106.809
27.3906 | -106.8046
27.3942 | -106.7997
27.3978 | -106.7958
27.4014 | -106.79
27.405 | -106.784
27.4086 | -106.7765
27.4123 | -106.7706
27.4159 | -106.7648
27.4195 | -106.757
27.4231 | -106.7485
27.4267 | -106.7405
27.4303 | -106.7321
27.4339 | -106.7225
27.4376 | -106.7117
27.4412 | -106.7027
27.4448 | -106.6943
27.4484 | -106.6835
27.452 | -106.6725
27.4556 | -106.6611
27.4592 | -106.6492
27.4628 | -106.6402
27.4665 | -106.6309
27.4701 | -106.621
27.4737 | -106.6121
27.4773 | -106.6029
27.4809 | -106.5947
27.4845 | -106.5852
27.4881 | -106.5745
27.4918 | -106.5641
27.4954 | -106.5547
27.499 | -106.5463
27.5026 | -106.539
27.5062 | -106.5323
27.5098 | -106.5248
27.5134 | -106.5186
27.5171 | -106.5149
27.5207 | -106.5109
27.5243 | -106.5057
27.5279 | -106.5007
27.5315 | -106.496
27.5351 | -106.4921
27.5387 | -106.4874
27.5423 | -106.4842
27.546 | -106.4828
27.5496 | -106.4815
27.5532 | -106.48
27.5568 | -106.4795
27.5604 | -106.4794
27.564 | -106.4802
27.5676 | -106.4802
27.5713 | -106.4806
27.5749 | -106.4822
27.5785 | -106.4846
27.5821 | -106.4869
27.5857 | -106.4915
27.5893 | -106.4951
27.5929 | -106.499
27.5965 | -106.5047
27.6002 | -106.5097
27.6038 | -106.5141
27.6074 | -106.5187
27.611 | -106.5242
27.6146 | -106.5307
27.6182 | -106.5374
27.6218 | -106.5435
27.6255 | -106.5505
27.6291 | -106.5553
27.6327 | -106.56
27.6363 | -106.5645
27.6399 | -106.5688
27.6435 | -106.5733
27.6471 | -106.5788
27.6508 | -106.5835
27.6544 | -106.5871
27.658 | -106.5915
27.6616 | -106.5959
27.6652 | -106.6003
27.6688 | -106.6044
27.6724 | -106.6065
27.676 | -106.6101
27.6797 | -106.6142
27.6833 | -106.6173
27.6869 | -106.6193
27.6905 | -106.6222
27.6941 | -106.6239
27.6977 | -106.6246
27.7013 | -106.6263
27.705 | -106.6287
27.7086 | -106.6306
27.7122 | -106.6312
27.7158 | -106.6324
27.7194 | -106.6348
27.723 | -106.6359
27.7266 | -106.6352
27.7303 | -106.6354
27.7339 | -106.6349
27.7375 | -106.6321
27.7411 | -106.6309
27.7447 | -106.6297
27.7483 | -106.6277
27.7519 | -106.6255
27.7555 | -106.625
27.7592 | -106.6239
27.7628 | -106.623
27.7664 | -106.6212
27.77 | -106.6205
27.7736 | -106.6199
27.7772 | -106.6188
27.7808 | -106.6173
27.7845 | -106.6151
27.7881 | -106.6111
27.7917 | -106.6088
27.7953 | -106.6065
27.7989 | -106.6045
27.8025 | -106.603
27.8061 | -106.6011
27.8097 | -106.6001
27.8134 | -106.5987
27.817 | -106.5966
27.8206 | -106.5967
27.8242 | -106.5962
27.8278 | -106.595
27.8314 | -106.5931
27.835 | -106.5912
27.8387 | -106.5889
27.8423 | -106.5878
27.8459 | -106.5865
27.8495 | -106.5848
27.8531 | -106.5826
27.8567 | -106.5807
27.8603 | -106.5788
27.8639 | -106.5774
27.8676 | -106.5764
27.8712 | -106.5756
27.8748 | -106.5748
27.8784 | -106.5749
27.882 | -106.5747
27.8856 | -106.5737
27.8892 | -106.5719
27.8929 | -106.569
27.8965 | -106.567
27.9001 | -106.5656
27.9037 | -106.563
27.9073 | -106.561
27.9109 | -106.5585
27.9145 | -106.5562
27.9182 | -106.5539
27.9218 | -106.5513
27.9254 | -106.5472
27.929 | -106.5439
27.9326 | -106.5406
27.9362 | -106.5352
27.9398 | -106.531
27.9434 | -106.5281
27.9471 | -106.5246
27.9507 | -106.5204
27.9543 | -106.5171
27.9579 | -106.5136
27.9615 | -106.5097
27.9651 | -106.5054
27.9687 | -106.5011
27.9724 | -106.497
27.976 | -106.4929
27.9796 | -106.4895
27.9832 | -106.4862
27.9868 | -106.4804
27.9904 | -106.4746
27.994 | -106.4716
27.9977 | -106.4684
28.0013 | -106.4643