## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 224 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 226
Hood | 02-92B/P | 165
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1014.4 | hPa |  | PRC PRO 01
Humidity | 38.8 | %/Hr
Internal temperature of calibration bench | 23.1 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1017 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 23.7 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.96 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.03 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3123 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 14/05/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 18/05/2018
Engineer | Anthony HUE | Approval |  | 18/05/2018
CEO | Robert GLEMOT | Approval |  | 18/05/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4211 | 57.6259
0.012209 | -35.9888 | 38.4234
0.018313 | -34.9248 | 28.2655
0.024417 | -34.4604 | 22.002
0.030521 | -34.3458 | 18.2298
0.036626 | -34.2014 | 15.1717
0.04273 | -34.106 | 12.8049
0.048834 | -34.059 | 11.2738
0.054938 | -33.9942 | 10.1246
0.061043 | -33.9342 | 9.0341
0.067147 | -33.9238 | 8.2102
0.073251 | -33.9241 | 7.4535
0.079355 | -33.9413 | 6.7963
0.08546 | -33.9549 | 6.2038
0.091564 | -33.9505 | 5.7595
0.097668 | -33.9326 | 5.5012
0.10377 | -33.9109 | 5.2472
0.10988 | -33.8942 | 4.9658
0.11598 | -33.8854 | 4.6725
0.12209 | -33.8871 | 4.3903
0.12819 | -33.8963 | 4.1102
0.13429 | -33.9063 | 3.8271
0.1404 | -33.9093 | 3.5559
0.1465 | -33.9113 | 3.3041
0.15261 | -33.9091 | 3.088
0.15871 | -33.9035 | 2.9254
0.16482 | -33.8964 | 2.8169
0.17092 | -33.8944 | 2.7664
0.17702 | -33.8926 | 2.6867
0.18313 | -33.8919 | 2.603
0.18923 | -33.8917 | 2.4929
0.19534 | -33.8906 | 2.3684
0.20144 | -33.8895 | 2.231
0.20754 | -33.8871 | 2.1152
0.21365 | -33.8864 | 2.004
0.21975 | -33.8862 | 1.9084
0.22586 | -33.8873 | 1.822
0.23196 | -33.8897 | 1.7491
0.23807 | -33.8936 | 1.6845
0.24417 | -33.8974 | 1.6242
0.25027 | -33.9011 | 1.5604
0.25638 | -33.9034 | 1.4979
0.26248 | -33.9051 | 1.4318
0.26859 | -33.9061 | 1.3672
0.27469 | -33.9065 | 1.2979
0.2808 | -33.9073 | 1.2257
0.2869 | -33.908 | 1.1563
0.293 | -33.9088 | 1.0905
0.29911 | -33.9097 | 1.0311
0.30521 | -33.9105 | 0.97307
0.31132 | -33.9113 | 0.93444
0.31742 | -33.9126 | 0.89142
0.32353 | -33.9131 | 0.85661
0.32963 | -33.9143 | 0.82727
0.33573 | -33.9157 | 0.80588
0.34184 | -33.9175 | 0.7833
0.34794 | -33.92 | 0.75725
0.35405 | -33.9225 | 0.73131
0.36015 | -33.9253 | 0.70374
0.36626 | -33.9276 | 0.67419
0.37236 | -33.9299 | 0.63535
0.37846 | -33.932 | 0.58914
0.38457 | -33.9324 | 0.53058
0.39067 | -33.9333 | 0.48254
0.39678 | -33.9336 | 0.42979
0.40288 | -33.9337 | 0.37468
0.40899 | -33.9333 | 0.32207
0.41509 | -33.9323 | 0.27141
0.42119 | -33.9321 | 0.22839
0.4273 | -33.9318 | 0.18185
0.4334 | -33.9315 | 0.13561
0.43951 | -33.9317 | 0.092065
0.44561 | -33.9326 | 0.051561
0.45172 | -33.9343 | 0.018812
0.45782 | -33.9352 | -0.019062
0.46392 | -33.9366 | -0.056689
0.47003 | -33.9378 | -0.089835
0.47613 | -33.9387 | -0.12537
0.48224 | -33.9393 | -0.16174
0.48834 | -33.9383 | -0.19839
0.49445 | -33.9381 | -0.23526
0.50055 | -33.9375 | -0.26912
0.50665 | -33.9367 | -0.30511
0.51276 | -33.9357 | -0.34052
0.51886 | -33.9345 | -0.37572
0.52497 | -33.9331 | -0.40019
0.53107 | -33.9323 | -0.43033
0.53717 | -33.9315 | -0.45942
0.54328 | -33.9307 | -0.48268
0.54938 | -33.9305 | -0.50506
0.55549 | -33.9306 | -0.52561
0.56159 | -33.9318 | -0.53556
0.5677 | -33.9322 | -0.5489
0.5738 | -33.9334 | -0.56398
0.5799 | -33.9343 | -0.57747
0.58601 | -33.9355 | -0.58591
0.59211 | -33.9373 | -0.59035
0.59822 | -33.9388 | -0.59799
0.60432 | -33.9404 | -0.60848
0.61043 | -33.9418 | -0.61486
0.61653 | -33.9431 | -0.62397
0.62263 | -33.9448 | -0.63182
0.62874 | -33.9454 | -0.64189
0.63484 | -33.9464 | -0.6503
0.64095 | -33.9469 | -0.65984
0.64705 | -33.9473 | -0.66927
0.65316 | -33.9475 | -0.68133
0.65926 | -33.9477 | -0.69727
0.66536 | -33.9466 | -0.71587
0.67147 | -33.9463 | -0.72954
0.67757 | -33.946 | -0.74438
0.68368 | -33.9453 | -0.76414
0.68978 | -33.9446 | -0.78248
0.69589 | -33.9439 | -0.80242
0.70199 | -33.9431 | -0.82637
0.70809 | -33.9424 | -0.84684
0.7142 | -33.9418 | -0.86825
0.7203 | -33.9409 | -0.88879
0.72641 | -33.9404 | -0.91227
0.73251 | -33.9399 | -0.93328
0.73862 | -33.9402 | -0.95528
0.74472 | -33.94 | -0.9762
0.75082 | -33.9398 | -0.9989
0.75693 | -33.9402 | -1.0211
0.76303 | -33.9409 | -1.0431
0.76914 | -33.9426 | -1.0623
0.77524 | -33.9435 | -1.0816
0.78135 | -33.9447 | -1.1007
0.78745 | -33.9461 | -1.1198
0.79355 | -33.9472 | -1.1386
0.79966 | -33.9488 | -1.1569
0.80576 | -33.9509 | -1.1748
0.81187 | -33.9524 | -1.1951
0.81797 | -33.9543 | -1.2146
0.82408 | -33.9559 | -1.2337
0.83018 | -33.9575 | -1.2526
0.83628 | -33.9592 | -1.2696
0.84239 | -33.9603 | -1.2887
0.84849 | -33.9617 | -1.3088
0.8546 | -33.9631 | -1.3273
0.8607 | -33.9642 | -1.3469
0.86681 | -33.9653 | -1.3649
0.87291 | -33.9662 | -1.3827
0.87901 | -33.9666 | -1.4005
0.88512 | -33.9674 | -1.4179
0.89122 | -33.9681 | -1.4354
0.89733 | -33.9685 | -1.4508
0.90343 | -33.9689 | -1.468
0.90953 | -33.9691 | -1.483
0.91564 | -33.9692 | -1.499
0.92174 | -33.9694 | -1.514
0.92785 | -33.9695 | -1.5288
0.93395 | -33.9697 | -1.5429
0.94006 | -33.9699 | -1.5565
0.94616 | -33.9697 | -1.5691
0.95226 | -33.9698 | -1.5825
0.95837 | -33.9699 | -1.5969
0.96447 | -33.9701 | -1.6104
0.97058 | -33.9701 | -1.6243
0.97668 | -33.9699 | -1.6384
0.98279 | -33.9697 | -1.6517
0.98889 | -33.9695 | -1.6632
0.99499 | -33.9695 | -1.677
1.0011 | -33.9695 | -1.6884
1.0072 | -33.9692 | -1.7007
1.0133 | -33.9689 | -1.7121
1.0194 | -33.9686 | -1.7252
1.0255 | -33.9683 | -1.7391
1.0316 | -33.9679 | -1.7514
1.0377 | -33.9675 | -1.7651
1.0438 | -33.967 | -1.7781
1.0499 | -33.9663 | -1.7931
1.056 | -33.9657 | -1.8075
1.0621 | -33.9652 | -1.8219
1.0682 | -33.9648 | -1.8343
1.0743 | -33.9642 | -1.8496
1.0805 | -33.9637 | -1.8639
1.0866 | -33.9629 | -1.8786
1.0927 | -33.9621 | -1.8928
1.0988 | -33.9616 | -1.9076
1.1049 | -33.9608 | -1.9216
1.111 | -33.9601 | -1.9366
1.1171 | -33.9595 | -1.9509
1.1232 | -33.9591 | -1.9657
1.1293 | -33.9587 | -1.9796
1.1354 | -33.9585 | -1.9943
1.1415 | -33.9581 | -2.0096
1.1476 | -33.9576 | -2.0248
1.1537 | -33.9571 | -2.0391
1.1598 | -33.9572 | -2.0538
1.1659 | -33.9571 | -2.0669
1.172 | -33.9568 | -2.0807
1.1781 | -33.9566 | -2.0945
1.1842 | -33.9565 | -2.1085
1.1903 | -33.9569 | -2.1195
1.1964 | -33.9569 | -2.1329
1.2025 | -33.9571 | -2.1467
1.2086 | -33.9574 | -2.1591
1.2147 | -33.9575 | -2.1716
1.2209 | -33.9579 | -2.1829
1.227 | -33.9587 | -2.1939
1.2331 | -33.9591 | -2.2058
1.2392 | -33.9595 | -2.2179
1.2453 | -33.96 | -2.2289
1.2514 | -33.9605 | -2.239
1.2575 | -33.9611 | -2.2502
1.2636 | -33.962 | -2.2595
1.2697 | -33.9627 | -2.2689
1.2758 | -33.9634 | -2.2802
1.2819 | -33.9642 | -2.2905
1.288 | -33.9648 | -2.3016
1.2941 | -33.9657 | -2.3103
1.3002 | -33.9665 | -2.3186
1.3063 | -33.967 | -2.3279
1.3124 | -33.9677 | -2.3371
1.3185 | -33.9685 | -2.3464
1.3246 | -33.9692 | -2.3558
1.3307 | -33.9701 | -2.3646
1.3368 | -33.971 | -2.3734
1.3429 | -33.9716 | -2.3825
1.349 | -33.9722 | -2.392
1.3551 | -33.9729 | -2.4017
1.3613 | -33.9735 | -2.4103
1.3674 | -33.9742 | -2.4204
1.3735 | -33.9747 | -2.4303
1.3796 | -33.9752 | -2.4396
1.3857 | -33.9755 | -2.449
1.3918 | -33.9761 | -2.4595
1.3979 | -33.9767 | -2.4687
1.404 | -33.9772 | -2.4799
1.4101 | -33.9777 | -2.4905
1.4162 | -33.9781 | -2.5013
1.4223 | -33.9786 | -2.5126
1.4284 | -33.9792 | -2.5226
1.4345 | -33.9794 | -2.5342
1.4406 | -33.9798 | -2.5477
1.4467 | -33.9803 | -2.5593
1.4528 | -33.9808 | -2.5706
1.4589 | -33.9813 | -2.5824
1.465 | -33.9816 | -2.5933
1.4711 | -33.9821 | -2.6045
1.4772 | -33.9827 | -2.6176
1.4833 | -33.9832 | -2.6306
1.4894 | -33.9838 | -2.6426
1.4955 | -33.9844 | -2.6544
1.5016 | -33.985 | -2.6667
1.5078 | -33.9854 | -2.6806
1.5139 | -33.9859 | -2.6935
1.52 | -33.9864 | -2.7061
1.5261 | -33.9868 | -2.7187
1.5322 | -33.9872 | -2.7326
1.5383 | -33.9876 | -2.7452
1.5444 | -33.9882 | -2.7579
1.5505 | -33.9888 | -2.7711
1.5566 | -33.9894 | -2.7842
1.5627 | -33.9899 | -2.7979
1.5688 | -33.9906 | -2.8113
1.5749 | -33.991 | -2.8239
1.581 | -33.9914 | -2.8373
1.5871 | -33.9919 | -2.8511
1.5932 | -33.9924 | -2.864
1.5993 | -33.9928 | -2.878
1.6054 | -33.9933 | -2.891
1.6115 | -33.9938 | -2.905
1.6176 | -33.9942 | -2.9176
1.6237 | -33.9945 | -2.931
1.6298 | -33.9948 | -2.9436
1.6359 | -33.9952 | -2.9562
1.642 | -33.9957 | -2.9682
1.6482 | -33.996 | -2.981
1.6543 | -33.9964 | -2.9945
1.6604 | -33.9967 | -3.0066
1.6665 | -33.9969 | -3.019
1.6726 | -33.9972 | -3.0316
1.6787 | -33.9976 | -3.0441
1.6848 | -33.9976 | -3.0564
1.6909 | -33.9979 | -3.069
1.697 | -33.9981 | -3.0813
1.7031 | -33.9982 | -3.0938
1.7092 | -33.9983 | -3.1059
1.7153 | -33.9985 | -3.1178
1.7214 | -33.9983 | -3.1297
1.7275 | -33.9983 | -3.142
1.7336 | -33.9983 | -3.1547
1.7397 | -33.9985 | -3.1669
1.7458 | -33.9985 | -3.179
1.7519 | -33.9985 | -3.191
1.758 | -33.9986 | -3.2028
1.7641 | -33.9988 | -3.2144
1.7702 | -33.9987 | -3.227
1.7763 | -33.9987 | -3.2382
1.7824 | -33.9987 | -3.2504
1.7885 | -33.9987 | -3.2623
1.7947 | -33.9989 | -3.2742
1.8008 | -33.9989 | -3.2852
1.8069 | -33.9989 | -3.2963
1.813 | -33.9991 | -3.3083
1.8191 | -33.999 | -3.3206
1.8252 | -33.9991 | -3.3329
1.8313 | -33.9992 | -3.3448
1.8374 | -33.9993 | -3.3564
1.8435 | -33.9994 | -3.3679
1.8496 | -33.9994 | -3.3807
1.8557 | -33.9996 | -3.3932
1.8618 | -33.9999 | -3.4053
1.8679 | -34 | -3.4168
1.874 | -34.0001 | -3.4284
1.8801 | -34.0002 | -3.4406
1.8862 | -34.0005 | -3.4523
1.8923 | -34.0006 | -3.4641
1.8984 | -34.0006 | -3.4758
1.9045 | -34.0007 | -3.4871
1.9106 | -34.0007 | -3.499
1.9167 | -34.0007 | -3.5109
1.9228 | -34.0009 | -3.5224
1.9289 | -34.0009 | -3.5341
1.9351 | -34.001 | -3.5445
1.9412 | -34.0013 | -3.5555
1.9473 | -34.0015 | -3.5674
1.9534 | -34.0017 | -3.5787
1.9595 | -34.0019 | -3.5904
1.9656 | -34.0021 | -3.603
1.9717 | -34.0023 | -3.6147
1.9778 | -34.0026 | -3.6259
1.9839 | -34.0028 | -3.6371
1.99 | -34.003 | -3.648
1.9961 | -34.003 | -3.6603
2.0022 | -34.0032 | -3.6719
2.0083 | -34.0032 | -3.6831
2.0144 | -34.0034 | -3.6945
2.0205 | -34.0037 | -3.7063
2.0266 | -34.0041 | -3.7177
2.0327 | -34.0043 | -3.7294
2.0388 | -34.0047 | -3.7411
2.0449 | -34.0049 | -3.7526
2.051 | -34.0051 | -3.764
2.0571 | -34.0054 | -3.7754
2.0632 | -34.0056 | -3.7871
2.0693 | -34.0058 | -3.7996
2.0754 | -34.006 | -3.8113
2.0816 | -34.0063 | -3.8236
2.0877 | -34.0065 | -3.8355
2.0938 | -34.0069 | -3.8477
2.0999 | -34.0071 | -3.8598
2.106 | -34.0073 | -3.8719
2.1121 | -34.0075 | -3.8843
2.1182 | -34.0077 | -3.8953
2.1243 | -34.0079 | -3.9064
2.1304 | -34.0081 | -3.9179
2.1365 | -34.0084 | -3.9299
2.1426 | -34.0087 | -3.9424
2.1487 | -34.009 | -3.9539
2.1548 | -34.0094 | -3.9658
2.1609 | -34.0099 | -3.9773
2.167 | -34.0103 | -3.9887
2.1731 | -34.0106 | -4.001
2.1792 | -34.0108 | -4.0126
2.1853 | -34.0113 | -4.0247
2.1914 | -34.0116 | -4.0364
2.1975 | -34.012 | -4.0482
2.2036 | -34.0122 | -4.0606
2.2097 | -34.0125 | -4.0725
2.2158 | -34.0129 | -4.0834
2.222 | -34.0132 | -4.0943
2.2281 | -34.0135 | -4.1061
2.2342 | -34.0138 | -4.1178
2.2403 | -34.014 | -4.129
2.2464 | -34.0144 | -4.14
2.2525 | -34.0147 | -4.1517
2.2586 | -34.015 | -4.1626
2.2647 | -34.0153 | -4.1735
2.2708 | -34.0157 | -4.185
2.2769 | -34.016 | -4.196
2.283 | -34.0163 | -4.2073
2.2891 | -34.0165 | -4.2185
2.2952 | -34.0168 | -4.2294
2.3013 | -34.017 | -4.2406
2.3074 | -34.0173 | -4.2514
2.3135 | -34.0176 | -4.2629
2.3196 | -34.0178 | -4.2734
2.3257 | -34.0181 | -4.2843
2.3318 | -34.0183 | -4.2956
2.3379 | -34.0185 | -4.3066
2.344 | -34.0187 | -4.318
2.3501 | -34.019 | -4.3289
2.3562 | -34.0191 | -4.3392
2.3623 | -34.0193 | -4.3498
2.3685 | -34.0195 | -4.3605
2.3746 | -34.0197 | -4.3716
2.3807 | -34.0199 | -4.3822
2.3868 | -34.0201 | -4.3928
2.3929 | -34.0204 | -4.4035
2.399 | -34.0205 | -4.4139
2.4051 | -34.0208 | -4.4246
2.4112 | -34.021 | -4.4352
2.4173 | -34.0212 | -4.4459
2.4234 | -34.0213 | -4.4563
2.4295 | -34.0215 | -4.467
2.4356 | -34.0217 | -4.4777
2.4417 | -34.0218 | -4.4886
2.4478 | -34.022 | -4.4992
2.4539 | -34.0222 | -4.5095
2.46 | -34.0223 | -4.5203
2.4661 | -34.0225 | -4.5312
2.4722 | -34.0227 | -4.5413
2.4783 | -34.0228 | -4.5514
2.4844 | -34.0229 | -4.5621
2.4905 | -34.023 | -4.5727
2.4966 | -34.0232 | -4.5831
2.5027 | -34.0234 | -4.5932
2.5089 | -34.0235 | -4.6037
2.515 | -34.0236 | -4.614
2.5211 | -34.0238 | -4.6243
2.5272 | -34.0239 | -4.6345
2.5333 | -34.024 | -4.6452
2.5394 | -34.0242 | -4.6556
2.5455 | -34.0243 | -4.6665
2.5516 | -34.0246 | -4.6765
2.5577 | -34.0248 | -4.6876
2.5638 | -34.0249 | -4.6987
2.5699 | -34.025 | -4.7095
2.576 | -34.0251 | -4.7205
2.5821 | -34.0252 | -4.7313
2.5882 | -34.0253 | -4.7419
2.5943 | -34.0255 | -4.7522
2.6004 | -34.0258 | -4.7626
2.6065 | -34.0259 | -4.7723
2.6126 | -34.0261 | -4.7822
2.6187 | -34.0263 | -4.792
2.6248 | -34.0265 | -4.8022
2.6309 | -34.0267 | -4.8132
2.637 | -34.027 | -4.8244
2.6431 | -34.0272 | -4.8342
2.6492 | -34.0273 | -4.845
2.6554 | -34.0276 | -4.8554
2.6615 | -34.0277 | -4.8657
2.6676 | -34.028 | -4.8767
2.6737 | -34.0284 | -4.8875
2.6798 | -34.0287 | -4.8981
2.6859 | -34.0289 | -4.9088
2.692 | -34.0291 | -4.9191
2.6981 | -34.0294 | -4.9298
2.7042 | -34.0296 | -4.9407
2.7103 | -34.0298 | -4.9509
2.7164 | -34.0302 | -4.9614
2.7225 | -34.0305 | -4.9723
2.7286 | -34.0308 | -4.983
2.7347 | -34.031 | -4.9931
2.7408 | -34.0313 | -5.0029
2.7469 | -34.0316 | -5.014
2.753 | -34.0319 | -5.0241
2.7591 | -34.0323 | -5.0351
2.7652 | -34.0325 | -5.0463
2.7713 | -34.0327 | -5.0568
2.7774 | -34.033 | -5.0677
2.7835 | -34.0332 | -5.0784
2.7896 | -34.0334 | -5.0886
2.7958 | -34.0338 | -5.099
2.8019 | -34.0341 | -5.109
2.808 | -34.0344 | -5.1195
2.8141 | -34.0348 | -5.1308
2.8202 | -34.0351 | -5.1414
2.8263 | -34.0355 | -5.1521
2.8324 | -34.0359 | -5.1633
2.8385 | -34.0363 | -5.174
2.8446 | -34.0367 | -5.1848
2.8507 | -34.0371 | -5.1955
2.8568 | -34.0375 | -5.2058
2.8629 | -34.0378 | -5.2165
2.869 | -34.0382 | -5.2276
2.8751 | -34.0385 | -5.2382
2.8812 | -34.039 | -5.2495
2.8873 | -34.0393 | -5.2603
2.8934 | -34.0397 | -5.2712
2.8995 | -34.04 | -5.2822
2.9056 | -34.0403 | -5.2934
2.9117 | -34.0407 | -5.3043
2.9178 | -34.0411 | -5.3153
2.9239 | -34.0415 | -5.3268
2.93 | -34.042 | -5.3381
2.9361 | -34.0423 | -5.349
2.9423 | -34.0427 | -5.3597
2.9484 | -34.0432 | -5.3706
2.9545 | -34.0436 | -5.3817
2.9606 | -34.044 | -5.3932
2.9667 | -34.0444 | -5.4043
2.9728 | -34.0447 | -5.4155
2.9789 | -34.0452 | -5.4265
2.985 | -34.0456 | -5.4379
2.9911 | -34.046 | -5.4493
2.9972 | -34.0465 | -5.4604
3.0033 | -34.0469 | -5.4718
3.0094 | -34.0474 | -5.4829
3.0155 | -34.0478 | -5.4939
3.0216 | -34.0482 | -5.5052
3.0277 | -34.0486 | -5.5169
3.0338 | -34.049 | -5.5284
3.0399 | -34.0494 | -5.5398
3.046 | -34.0498 | -5.5514
3.0521 | -34.0501 | -5.5625
3.0582 | -34.0504 | -5.5744
3.0643 | -34.0508 | -5.5858
3.0704 | -34.0511 | -5.5977
3.0765 | -34.0514 | -5.6088
3.0827 | -34.0518 | -5.6202
3.0888 | -34.0521 | -5.6316
3.0949 | -34.0525 | -5.643
3.101 | -34.0529 | -5.6544
3.1071 | -34.0532 | -5.666
3.1132 | -34.0536 | -5.6776
3.1193 | -34.0539 | -5.6892
3.1254 | -34.0543 | -5.7005
3.1315 | -34.0547 | -5.7123
3.1376 | -34.055 | -5.7237
3.1437 | -34.0554 | -5.7352
3.1498 | -34.0557 | -5.7471
3.1559 | -34.056 | -5.7585
3.162 | -34.0563 | -5.7704
3.1681 | -34.0566 | -5.7823
3.1742 | -34.0569 | -5.7938
3.1803 | -34.0572 | -5.8053
3.1864 | -34.0576 | -5.8166
3.1925 | -34.0579 | -5.8279
3.1986 | -34.0581 | -5.8397
3.2047 | -34.0584 | -5.8517
3.2108 | -34.0587 | -5.8627
3.2169 | -34.059 | -5.8743
3.223 | -34.0592 | -5.8859
3.2292 | -34.0595 | -5.8974
3.2353 | -34.0596 | -5.908
3.2414 | -34.0599 | -5.9194
3.2475 | -34.0601 | -5.9306
3.2536 | -34.0604 | -5.942
3.2597 | -34.0606 | -5.9528
3.2658 | -34.0608 | -5.9637
3.2719 | -34.061 | -5.9745
3.278 | -34.0612 | -5.9859
3.2841 | -34.0615 | -5.9969
3.2902 | -34.0618 | -6.0082
3.2963 | -34.062 | -6.0196
3.3024 | -34.062 | -6.0304
3.3085 | -34.0623 | -6.0418
3.3146 | -34.0625 | -6.0528
3.3207 | -34.0627 | -6.0641
3.3268 | -34.0629 | -6.0754
3.3329 | -34.0631 | -6.0861
3.339 | -34.0632 | -6.0967
3.3451 | -34.0634 | -6.1079
3.3512 | -34.0635 | -6.1185
3.3573 | -34.0638 | -6.1294
3.3634 | -34.0639 | -6.1405
3.3696 | -34.0641 | -6.1518
3.3757 | -34.0642 | -6.1624
3.3818 | -34.0644 | -6.1731
3.3879 | -34.0646 | -6.1842
3.394 | -34.0647 | -6.1953
3.4001 | -34.0649 | -6.206
3.4062 | -34.065 | -6.2167
3.4123 | -34.0651 | -6.2271
3.4184 | -34.0653 | -6.2382
3.4245 | -34.0655 | -6.2488
3.4306 | -34.0655 | -6.2593
3.4367 | -34.0657 | -6.2697
3.4428 | -34.0658 | -6.2799
3.4489 | -34.066 | -6.2907
3.455 | -34.0661 | -6.3015
3.4611 | -34.0663 | -6.3126
3.4672 | -34.0664 | -6.3233
3.4733 | -34.0665 | -6.3343
3.4794 | -34.0667 | -6.345
3.4855 | -34.0668 | -6.3553
3.4916 | -34.067 | -6.3655
3.4977 | -34.0671 | -6.3766
3.5038 | -34.0672 | -6.3874
3.5099 | -34.0674 | -6.3975
3.5161 | -34.0675 | -6.4078
3.5222 | -34.0676 | -6.4181
3.5283 | -34.0677 | -6.4287
3.5344 | -34.0678 | -6.4394
3.5405 | -34.068 | -6.4498
3.5466 | -34.0681 | -6.4606
3.5527 | -34.0682 | -6.4712
3.5588 | -34.0683 | -6.4817
3.5649 | -34.0683 | -6.4921
3.571 | -34.0684 | -6.5026
3.5771 | -34.0685 | -6.5136
3.5832 | -34.0687 | -6.5234
3.5893 | -34.0689 | -6.534
3.5954 | -34.0691 | -6.5447
3.6015 | -34.0692 | -6.5554
3.6076 | -34.0694 | -6.5659
3.6137 | -34.0696 | -6.5759
3.6198 | -34.0698 | -6.586
3.6259 | -34.0699 | -6.5968
3.632 | -34.0701 | -6.6073
3.6381 | -34.0703 | -6.618
3.6442 | -34.0704 | -6.6283
3.6503 | -34.0706 | -6.6384
3.6565 | -34.0708 | -6.6491
3.6626 | -34.071 | -6.6595
3.6687 | -34.0712 | -6.6702
3.6748 | -34.0715 | -6.6805
3.6809 | -34.0717 | -6.6911
3.687 | -34.0719 | -6.7019
3.6931 | -34.0721 | -6.7121
3.6992 | -34.0724 | -6.7227
3.7053 | -34.0725 | -6.7331
3.7114 | -34.0727 | -6.7437
3.7175 | -34.0729 | -6.7544
3.7236 | -34.0732 | -6.7644
3.7297 | -34.0734 | -6.7752
3.7358 | -34.0736 | -6.786
3.7419 | -34.0738 | -6.7967
3.748 | -34.074 | -6.8071
3.7541 | -34.0742 | -6.8176
3.7602 | -34.0745 | -6.828
3.7663 | -34.0747 | -6.8383
3.7724 | -34.075 | -6.849
3.7785 | -34.0753 | -6.8593
3.7846 | -34.0755 | -6.8696
3.7907 | -34.0758 | -6.8806
3.7969 | -34.0761 | -6.8912
3.803 | -34.0763 | -6.9014
3.8091 | -34.0766 | -6.9119
3.8152 | -34.0769 | -6.9223
3.8213 | -34.0771 | -6.9331
3.8274 | -34.0774 | -6.9433
3.8335 | -34.0777 | -6.954
3.8396 | -34.078 | -6.9647
3.8457 | -34.0784 | -6.9753
3.8518 | -34.0787 | -6.9858
3.8579 | -34.0791 | -6.9962
3.864 | -34.0794 | -7.0069
3.8701 | -34.0797 | -7.0179
3.8762 | -34.08 | -7.0283
3.8823 | -34.0802 | -7.0391
3.8884 | -34.0806 | -7.0496
3.8945 | -34.081 | -7.0605
3.9006 | -34.0813 | -7.0714
3.9067 | -34.0816 | -7.0825
3.9128 | -34.0819 | -7.0932
3.9189 | -34.0823 | -7.1036
3.925 | -34.0825 | -7.1144
3.9311 | -34.0828 | -7.1245
3.9372 | -34.0832 | -7.1354
3.9434 | -34.0835 | -7.146
3.9495 | -34.0839 | -7.1567
3.9556 | -34.0842 | -7.1671
3.9617 | -34.0845 | -7.1774
3.9678 | -34.0849 | -7.1877
3.9739 | -34.0852 | -7.1983
3.98 | -34.0855 | -7.2091
3.9861 | -34.0858 | -7.2197
3.9922 | -34.0861 | -7.2304
3.9983 | -34.0864 | -7.2407
4.0044 | -34.0867 | -7.2509
4.0105 | -34.087 | -7.2615
4.0166 | -34.0874 | -7.2722
4.0227 | -34.0877 | -7.2828
4.0288 | -34.088 | -7.2939
4.0349 | -34.0882 | -7.3043
4.041 | -34.0885 | -7.3149
4.0471 | -34.0888 | -7.3257
4.0532 | -34.0891 | -7.3359
4.0593 | -34.0895 | -7.3467
4.0654 | -34.0898 | -7.3571
4.0715 | -34.0901 | -7.3682
4.0776 | -34.0905 | -7.3787
4.0838 | -34.0908 | -7.3897
4.0899 | -34.0911 | -7.4
4.096 | -34.0914 | -7.4106
4.1021 | -34.0917 | -7.421
4.1082 | -34.092 | -7.4314
4.1143 | -34.0923 | -7.4417
4.1204 | -34.0926 | -7.452
4.1265 | -34.0929 | -7.4623
4.1326 | -34.0932 | -7.4726
4.1387 | -34.0936 | -7.483
4.1448 | -34.0939 | -7.4936
4.1509 | -34.0941 | -7.5042
4.157 | -34.0944 | -7.5152
4.1631 | -34.0947 | -7.5258
4.1692 | -34.095 | -7.5363
4.1753 | -34.0953 | -7.5468
4.1814 | -34.0955 | -7.5569
4.1875 | -34.0959 | -7.5677
4.1936 | -34.0961 | -7.5783
4.1997 | -34.0964 | -7.5885
4.2058 | -34.0967 | -7.5992
4.2119 | -34.097 | -7.6095
4.218 | -34.0972 | -7.6197
4.2241 | -34.0975 | -7.63
4.2303 | -34.0978 | -7.6403
4.2364 | -34.0981 | -7.6508
4.2425 | -34.0983 | -7.6614
4.2486 | -34.0985 | -7.672
4.2547 | -34.0987 | -7.6822
4.2608 | -34.0989 | -7.6928
4.2669 | -34.0992 | -7.7028
4.273 | -34.0995 | -7.7134
4.2791 | -34.0997 | -7.7239
4.2852 | -34.1 | -7.7344
4.2913 | -34.1001 | -7.7446
4.2974 | -34.1004 | -7.7552
4.3035 | -34.1007 | -7.7659
4.3096 | -34.1009 | -7.7766
4.3157 | -34.1012 | -7.7868
4.3218 | -34.1014 | -7.7972
4.3279 | -34.1017 | -7.8075
4.334 | -34.1019 | -7.8178
4.3401 | -34.1021 | -7.8279
4.3462 | -34.1024 | -7.8383
4.3523 | -34.1027 | -7.8483
4.3584 | -34.1028 | -7.8583
4.3645 | -34.1031 | -7.8687
4.3707 | -34.1033 | -7.8788
4.3768 | -34.1035 | -7.8891
4.3829 | -34.1038 | -7.8998
4.389 | -34.104 | -7.9101
4.3951 | -34.1042 | -7.9202
4.4012 | -34.1044 | -7.9308
4.4073 | -34.1046 | -7.941
4.4134 | -34.1049 | -7.9511
4.4195 | -34.1051 | -7.9614
4.4256 | -34.1054 | -7.9715
4.4317 | -34.1057 | -7.9818
4.4378 | -34.1059 | -7.9918
4.4439 | -34.1061 | -8.0018
4.45 | -34.1062 | -8.012
4.4561 | -34.1064 | -8.0223
4.4622 | -34.1066 | -8.0326
4.4683 | -34.1069 | -8.0428
4.4744 | -34.1071 | -8.0534
4.4805 | -34.1073 | -8.0633
4.4866 | -34.1075 | -8.0735
4.4927 | -34.1078 | -8.084
4.4988 | -34.108 | -8.0942
4.5049 | -34.1082 | -8.104
4.511 | -34.1085 | -8.114
4.5172 | -34.1087 | -8.1242
4.5233 | -34.1089 | -8.1348
4.5294 | -34.1091 | -8.1449
4.5355 | -34.1093 | -8.1551
4.5416 | -34.1095 | -8.1653
4.5477 | -34.1098 | -8.1755
4.5538 | -34.11 | -8.1853
4.5599 | -34.1102 | -8.1955
4.566 | -34.1104 | -8.2056
4.5721 | -34.1106 | -8.2159
4.5782 | -34.1109 | -8.2263
4.5843 | -34.1111 | -8.2366
4.5904 | -34.1114 | -8.2469
4.5965 | -34.1116 | -8.257
4.6026 | -34.1118 | -8.2677
4.6087 | -34.112 | -8.2778
4.6148 | -34.1122 | -8.2883
4.6209 | -34.1125 | -8.2985
4.627 | -34.1127 | -8.3089
4.6331 | -34.1129 | -8.319
4.6392 | -34.1131 | -8.3293
4.6453 | -34.1134 | -8.3395
4.6514 | -34.1136 | -8.3498
4.6576 | -34.1138 | -8.3604
4.6637 | -34.114 | -8.3702
4.6698 | -34.1142 | -8.3802
4.6759 | -34.1144 | -8.3908
4.682 | -34.1146 | -8.4008
4.6881 | -34.1147 | -8.411
4.6942 | -34.115 | -8.4211
4.7003 | -34.1152 | -8.4314
4.7064 | -34.1154 | -8.4414
4.7125 | -34.1157 | -8.4519
4.7186 | -34.1159 | -8.4618
4.7247 | -34.1161 | -8.4721
4.7308 | -34.1164 | -8.4822
4.7369 | -34.1166 | -8.4925
4.743 | -34.1168 | -8.5031
4.7491 | -34.117 | -8.5136
4.7552 | -34.1172 | -8.524
4.7613 | -34.1174 | -8.5343
4.7674 | -34.1176 | -8.5443
4.7735 | -34.1178 | -8.555
4.7796 | -34.1181 | -8.5655
4.7857 | -34.1183 | -8.5759
4.7918 | -34.1186 | -8.586
4.7979 | -34.1188 | -8.5961
4.8041 | -34.119 | -8.6064
4.8102 | -34.1193 | -8.6163
4.8163 | -34.1195 | -8.6264
4.8224 | -34.1197 | -8.6367
4.8285 | -34.12 | -8.6466
4.8346 | -34.1202 | -8.6566
4.8407 | -34.1204 | -8.6669
4.8468 | -34.1207 | -8.6774
4.8529 | -34.1209 | -8.688
4.859 | -34.1212 | -8.6986
4.8651 | -34.1214 | -8.7085
4.8712 | -34.1216 | -8.7185
4.8773 | -34.1218 | -8.7289
4.8834 | -34.1221 | -8.7391
4.8895 | -34.1223 | -8.7491
4.8956 | -34.1225 | -8.7593
4.9017 | -34.1227 | -8.7694
4.9078 | -34.1229 | -8.78
4.9139 | -34.1232 | -8.7898
4.92 | -34.1234 | -8.7996
4.9261 | -34.1237 | -8.8095
4.9322 | -34.1239 | -8.8197
4.9383 | -34.1242 | -8.8299
4.9445 | -34.1244 | -8.8402
4.9506 | -34.1247 | -8.8502
4.9567 | -34.1249 | -8.8602
4.9628 | -34.1251 | -8.8703
4.9689 | -34.1254 | -8.8809
4.975 | -34.1256 | -8.8914
4.9811 | -34.1259 | -8.9014
4.9872 | -34.1262 | -8.9114
4.9933 | -34.1265 | -8.9219
4.9994 | -34.1267 | -8.9324
5.0055 | -34.127 | -8.9427
5.0116 | -34.1273 | -8.9533
5.0177 | -34.1276 | -8.9638
5.0238 | -34.1279 | -8.9742
5.0299 | -34.1282 | -8.9847
5.036 | -34.1285 | -8.995
5.0421 | -34.1287 | -9.0052
5.0482 | -34.129 | -9.0154
5.0543 | -34.1292 | -9.0254
5.0604 | -34.1295 | -9.0356
5.0665 | -34.1298 | -9.0456
5.0726 | -34.1301 | -9.0561
5.0787 | -34.1304 | -9.0665
5.0848 | -34.1307 | -9.0768
5.091 | -34.1309 | -9.0869
5.0971 | -34.1312 | -9.0968
5.1032 | -34.1315 | -9.1073
5.1093 | -34.1318 | -9.1176
5.1154 | -34.132 | -9.1276
5.1215 | -34.1323 | -9.1376
5.1276 | -34.1326 | -9.1477
5.1337 | -34.1329 | -9.1576
5.1398 | -34.1331 | -9.1676
5.1459 | -34.1334 | -9.178
5.152 | -34.1337 | -9.1884
5.1581 | -34.134 | -9.1988
5.1642 | -34.1343 | -9.2091
5.1703 | -34.1346 | -9.2193
5.1764 | -34.1349 | -9.2294
5.1825 | -34.1352 | -9.2393
5.1886 | -34.1355 | -9.2493
5.1947 | -34.1357 | -9.2595
5.2008 | -34.1361 | -9.2699
5.2069 | -34.1363 | -9.2803
5.213 | -34.1366 | -9.2903
5.2191 | -34.137 | -9.3005
5.2252 | -34.1372 | -9.3108
5.2314 | -34.1375 | -9.3212
5.2375 | -34.1378 | -9.3312
5.2436 | -34.1381 | -9.3414
5.2497 | -34.1384 | -9.3514
5.2558 | -34.1387 | -9.3619
5.2619 | -34.139 | -9.3724
5.268 | -34.1393 | -9.3828
5.2741 | -34.1396 | -9.3931
5.2802 | -34.1399 | -9.4037
5.2863 | -34.1403 | -9.4138
5.2924 | -34.1406 | -9.4243
5.2985 | -34.1409 | -9.4346
5.3046 | -34.1412 | -9.4449
5.3107 | -34.1415 | -9.4553
5.3168 | -34.1418 | -9.4656
5.3229 | -34.1421 | -9.4759
5.329 | -34.1424 | -9.4862
5.3351 | -34.1427 | -9.4962
5.3412 | -34.143 | -9.5064
5.3473 | -34.1433 | -9.5167
5.3534 | -34.1437 | -9.5271
5.3595 | -34.1439 | -9.5374
5.3656 | -34.1443 | -9.5477
5.3717 | -34.1446 | -9.5583
5.3779 | -34.1449 | -9.5689
5.384 | -34.1452 | -9.5792
5.3901 | -34.1455 | -9.5897
5.3962 | -34.1458 | -9.5998
5.4023 | -34.1462 | -9.6103
5.4084 | -34.1465 | -9.6206
5.4145 | -34.1469 | -9.6306
5.4206 | -34.1471 | -9.6409
5.4267 | -34.1475 | -9.6515
5.4328 | -34.1478 | -9.6619
5.4389 | -34.1481 | -9.672
5.445 | -34.1484 | -9.6826
5.4511 | -34.1487 | -9.6927
5.4572 | -34.1491 | -9.7031
5.4633 | -34.1494 | -9.7134
5.4694 | -34.1497 | -9.7237
5.4755 | -34.15 | -9.7338
5.4816 | -34.1503 | -9.7441
5.4877 | -34.1506 | -9.7543
5.4938 | -34.1509 | -9.7643
5.4999 | -34.1513 | -9.7745
5.506 | -34.1516 | -9.7848
5.5121 | -34.1519 | -9.795
5.5183 | -34.1523 | -9.8051
5.5244 | -34.1525 | -9.8156
5.5305 | -34.1529 | -9.8259
5.5366 | -34.1532 | -9.8359
5.5427 | -34.1535 | -9.8464
5.5488 | -34.1538 | -9.8567
5.5549 | -34.1541 | -9.8671
5.561 | -34.1544 | -9.8776
5.5671 | -34.1548 | -9.8877
5.5732 | -34.1551 | -9.8983
5.5793 | -34.1554 | -9.909
5.5854 | -34.1558 | -9.9192
5.5915 | -34.1561 | -9.9298
5.5976 | -34.1564 | -9.9402
5.6037 | -34.1567 | -9.9506
5.6098 | -34.157 | -9.9609
5.6159 | -34.1573 | -9.971
5.622 | -34.1576 | -9.9814
5.6281 | -34.158 | -9.9918
5.6342 | -34.1583 | -10.0024
5.6403 | -34.1586 | -10.0128
5.6464 | -34.1589 | -10.0233
5.6525 | -34.1592 | -10.0337
5.6586 | -34.1596 | -10.0437
5.6648 | -34.1599 | -10.0541
5.6709 | -34.1603 | -10.0646
5.677 | -34.1606 | -10.075
5.6831 | -34.1609 | -10.0854
5.6892 | -34.1612 | -10.0959
5.6953 | -34.1616 | -10.1062
5.7014 | -34.1619 | -10.1167
5.7075 | -34.1622 | -10.127
5.7136 | -34.1625 | -10.1373
5.7197 | -34.1628 | -10.1476
5.7258 | -34.1631 | -10.1577
5.7319 | -34.1634 | -10.1681
5.738 | -34.1638 | -10.1785
5.7441 | -34.1641 | -10.1888
5.7502 | -34.1645 | -10.1988
5.7563 | -34.1648 | -10.2092
5.7624 | -34.1652 | -10.2197
5.7685 | -34.1655 | -10.2301
5.7746 | -34.1658 | -10.2407
5.7807 | -34.166 | -10.2511
5.7868 | -34.1663 | -10.2616
5.7929 | -34.1666 | -10.2719
5.799 | -34.167 | -10.2821
5.8052 | -34.1673 | -10.2923
5.8113 | -34.1676 | -10.3025
5.8174 | -34.1679 | -10.3127
5.8235 | -34.1682 | -10.323
5.8296 | -34.1685 | -10.3336
5.8357 | -34.1688 | -10.3439
5.8418 | -34.1691 | -10.3543
5.8479 | -34.1694 | -10.3648
5.854 | -34.1697 | -10.3752
5.8601 | -34.1701 | -10.3854
5.8662 | -34.1704 | -10.3959
5.8723 | -34.1707 | -10.4061
5.8784 | -34.171 | -10.4163
5.8845 | -34.1713 | -10.4265
5.8906 | -34.1716 | -10.4369
5.8967 | -34.1719 | -10.447
5.9028 | -34.1722 | -10.4572
5.9089 | -34.1726 | -10.4674
5.915 | -34.1729 | -10.4775
5.9211 | -34.1732 | -10.4877
5.9272 | -34.1735 | -10.4978
5.9333 | -34.1738 | -10.508
5.9394 | -34.1742 | -10.5183
5.9455 | -34.1745 | -10.5284
5.9517 | -34.1748 | -10.5384
5.9578 | -34.1752 | -10.5489
5.9639 | -34.1755 | -10.5588
5.97 | -34.1759 | -10.5691
5.9761 | -34.1762 | -10.5793
5.9822 | -34.1765 | -10.5897
5.9883 | -34.1768 | -10.6
5.9944 | -34.1772 | -10.6103
6.0005 | -34.1775 | -10.6208
6.0066 | -34.1778 | -10.6308
6.0127 | -34.1782 | -10.6408
6.0188 | -34.1785 | -10.6514
6.0249 | -34.1788 | -10.6617
6.031 | -34.1791 | -10.6718
6.0371 | -34.1794 | -10.6821
6.0432 | -34.1797 | -10.6922
6.0493 | -34.18 | -10.7022
6.0554 | -34.1804 | -10.7123
6.0615 | -34.1807 | -10.7225
6.0676 | -34.1811 | -10.733
6.0737 | -34.1814 | -10.7433
6.0798 | -34.1817 | -10.7535
6.0859 | -34.1821 | -10.7637
6.0921 | -34.1824 | -10.7737
6.0982 | -34.1827 | -10.7839
6.1043 | -34.1831 | -10.7939
6.1104 | -34.1834 | -10.8044
6.1165 | -34.1837 | -10.8141
6.1226 | -34.1841 | -10.8243
6.1287 | -34.1844 | -10.8344
6.1348 | -34.1847 | -10.8445
6.1409 | -34.185 | -10.8547
6.147 | -34.1853 | -10.8648
6.1531 | -34.1856 | -10.8749
6.1592 | -34.186 | -10.885
6.1653 | -34.1863 | -10.895
6.1714 | -34.1866 | -10.9052
6.1775 | -34.1869 | -10.9155
6.1836 | -34.1872 | -10.9257
6.1897 | -34.1876 | -10.9359
6.1958 | -34.1879 | -10.9463
6.2019 | -34.1883 | -10.9566
6.208 | -34.1886 | -10.9668
6.2141 | -34.1889 | -10.9768
6.2202 | -34.1893 | -10.9871
6.2263 | -34.1896 | -10.9971
6.2325 | -34.1899 | -11.0069
6.2386 | -34.1902 | -11.0168
6.2447 | -34.1906 | -11.0272
6.2508 | -34.1909 | -11.0374
6.2569 | -34.1913 | -11.0474
6.263 | -34.1916 | -11.0576
6.2691 | -34.1919 | -11.0676
6.2752 | -34.1922 | -11.0778
6.2813 | -34.1925 | -11.0879
6.2874 | -34.1928 | -11.0979
6.2935 | -34.1931 | -11.108
6.2996 | -34.1934 | -11.1181
6.3057 | -34.1938 | -11.1283
6.3118 | -34.194 | -11.1385
6.3179 | -34.1944 | -11.1486
6.324 | -34.1947 | -11.1586
6.3301 | -34.195 | -11.1687
6.3362 | -34.1954 | -11.1787
6.3423 | -34.1957 | -11.1889
6.3484 | -34.196 | -11.1991
6.3545 | -34.1963 | -11.2089
6.3606 | -34.1966 | -11.2187
6.3667 | -34.1969 | -11.2286
6.3728 | -34.1973 | -11.2388
6.379 | -34.1976 | -11.2488
6.3851 | -34.1979 | -11.2587
6.3912 | -34.1982 | -11.2687
6.3973 | -34.1985 | -11.2788
6.4034 | -34.1988 | -11.2887
6.4095 | -34.1991 | -11.2987
6.4156 | -34.1995 | -11.3088
6.4217 | -34.1998 | -11.3188
6.4278 | -34.2001 | -11.3289
6.4339 | -34.2004 | -11.3389
6.44 | -34.2008 | -11.3488
6.4461 | -34.2011 | -11.3591
6.4522 | -34.2014 | -11.369
6.4583 | -34.2017 | -11.3791
6.4644 | -34.202 | -11.3893
6.4705 | -34.2023 | -11.399
6.4766 | -34.2026 | -11.409
6.4827 | -34.2029 | -11.4191
6.4888 | -34.2032 | -11.4289
6.4949 | -34.2035 | -11.4387
6.501 | -34.2039 | -11.4488
6.5071 | -34.2042 | -11.4585
6.5132 | -34.2045 | -11.4686
6.5194 | -34.2048 | -11.4787
6.5255 | -34.2051 | -11.4885
6.5316 | -34.2054 | -11.4986
6.5377 | -34.2057 | -11.5085
6.5438 | -34.206 | -11.5183
6.5499 | -34.2064 | -11.5282
6.556 | -34.2067 | -11.5382
6.5621 | -34.207 | -11.5484
6.5682 | -34.2073 | -11.5585
6.5743 | -34.2077 | -11.5684
6.5804 | -34.208 | -11.5784
6.5865 | -34.2083 | -11.5884
6.5926 | -34.2086 | -11.5983
6.5987 | -34.2089 | -11.608
6.6048 | -34.2092 | -11.618
6.6109 | -34.2096 | -11.6279
6.617 | -34.2098 | -11.638
6.6231 | -34.2101 | -11.648
6.6292 | -34.2105 | -11.6579
6.6353 | -34.2108 | -11.6682
6.6414 | -34.2111 | -11.6781
6.6475 | -34.2115 | -11.688
6.6536 | -34.2117 | -11.6978
6.6597 | -34.2121 | -11.7075
6.6659 | -34.2124 | -11.7175
6.672 | -34.2127 | -11.7273
6.6781 | -34.213 | -11.7373
6.6842 | -34.2134 | -11.7471
6.6903 | -34.2137 | -11.757
6.6964 | -34.214 | -11.767
6.7025 | -34.2143 | -11.7768
6.7086 | -34.2147 | -11.7867
6.7147 | -34.215 | -11.7967
6.7208 | -34.2153 | -11.8067
6.7269 | -34.2156 | -11.8166
6.733 | -34.2159 | -11.8262
6.7391 | -34.2162 | -11.8361
6.7452 | -34.2166 | -11.8458
6.7513 | -34.2169 | -11.8555
6.7574 | -34.2172 | -11.8655
6.7635 | -34.2175 | -11.8752
6.7696 | -34.2178 | -11.8849
6.7757 | -34.2182 | -11.8949
6.7818 | -34.2185 | -11.9048
6.7879 | -34.2189 | -11.9145
6.794 | -34.2192 | -11.9242
6.8001 | -34.2195 | -11.934
6.8063 | -34.2198 | -11.9438
6.8124 | -34.2201 | -11.9538
6.8185 | -34.2204 | -11.9638
6.8246 | -34.2207 | -11.9735
6.8307 | -34.221 | -11.9832
6.8368 | -34.2213 | -11.9933
6.8429 | -34.2217 | -12.0032
6.849 | -34.222 | -12.0131
6.8551 | -34.2224 | -12.0229
6.8612 | -34.2227 | -12.0327
6.8673 | -34.223 | -12.0425
6.8734 | -34.2233 | -12.0525
6.8795 | -34.2236 | -12.0624
6.8856 | -34.2239 | -12.0722
6.8917 | -34.2243 | -12.082
6.8978 | -34.2246 | -12.0916
6.9039 | -34.2249 | -12.1011
6.91 | -34.2252 | -12.111
6.9161 | -34.2256 | -12.1206
6.9222 | -34.2259 | -12.1304
6.9283 | -34.2262 | -12.1399
6.9344 | -34.2266 | -12.1499
6.9405 | -34.2269 | -12.1596
6.9466 | -34.2273 | -12.1696
6.9528 | -34.2276 | -12.1796
6.9589 | -34.228 | -12.1896
6.965 | -34.2283 | -12.1995
6.9711 | -34.2287 | -12.2094
6.9772 | -34.229 | -12.2192
6.9833 | -34.2293 | -12.229
6.9894 | -34.2297 | -12.239
6.9955 | -34.23 | -12.2486
7.0016 | -34.2303 | -12.2584
7.0077 | -34.2307 | -12.2682
7.0138 | -34.231 | -12.2779
7.0199 | -34.2313 | -12.2876
7.026 | -34.2317 | -12.2974
7.0321 | -34.232 | -12.307
7.0382 | -34.2323 | -12.3164
7.0443 | -34.2327 | -12.3264
7.0504 | -34.233 | -12.3363
7.0565 | -34.2334 | -12.3461
7.0626 | -34.2337 | -12.3556
7.0687 | -34.234 | -12.365
7.0748 | -34.2344 | -12.3745
7.0809 | -34.2347 | -12.3842
7.087 | -34.2351 | -12.394
7.0932 | -34.2354 | -12.4038
7.0993 | -34.2357 | -12.4135
7.1054 | -34.2361 | -12.4231
7.1115 | -34.2364 | -12.4327
7.1176 | -34.2368 | -12.4423
7.1237 | -34.2371 | -12.452
7.1298 | -34.2374 | -12.4617
7.1359 | -34.2377 | -12.4714
7.142 | -34.2381 | -12.4809
7.1481 | -34.2384 | -12.4905
7.1542 | -34.2387 | -12.5002
7.1603 | -34.2391 | -12.5099
7.1664 | -34.2394 | -12.5196
7.1725 | -34.2398 | -12.5294
7.1786 | -34.2402 | -12.539
7.1847 | -34.2405 | -12.5488
7.1908 | -34.2408 | -12.5587
7.1969 | -34.2411 | -12.5685
7.203 | -34.2415 | -12.5781
7.2091 | -34.2419 | -12.5877
7.2152 | -34.2422 | -12.5974
7.2213 | -34.2425 | -12.607
7.2274 | -34.2429 | -12.6166
7.2335 | -34.2432 | -12.6263
7.2397 | -34.2436 | -12.6359
7.2458 | -34.244 | -12.6454
7.2519 | -34.2443 | -12.6551
7.258 | -34.2446 | -12.6649
7.2641 | -34.245 | -12.6746
7.2702 | -34.2453 | -12.6845
7.2763 | -34.2457 | -12.6942
7.2824 | -34.2461 | -12.704
7.2885 | -34.2464 | -12.7137
7.2946 | -34.2467 | -12.7234
7.3007 | -34.2471 | -12.7332
7.3068 | -34.2474 | -12.7431
7.3129 | -34.2477 | -12.753
7.319 | -34.2481 | -12.7627
7.3251 | -34.2484 | -12.7722
7.3312 | -34.2488 | -12.7819
7.3373 | -34.2492 | -12.7914
7.3434 | -34.2495 | -12.801
7.3495 | -34.2499 | -12.8106
7.3556 | -34.2502 | -12.8204
7.3617 | -34.2506 | -12.8302
7.3678 | -34.2509 | -12.8398
7.3739 | -34.2513 | -12.8497
7.3801 | -34.2516 | -12.8596
7.3862 | -34.252 | -12.8693
7.3923 | -34.2523 | -12.8791
7.3984 | -34.2527 | -12.8886
7.4045 | -34.2531 | -12.8982
7.4106 | -34.2535 | -12.9079
7.4167 | -34.2538 | -12.9176
7.4228 | -34.2542 | -12.9274
7.4289 | -34.2545 | -12.9371
7.435 | -34.2549 | -12.9466
7.4411 | -34.2552 | -12.9564
7.4472 | -34.2556 | -12.9661
7.4533 | -34.2559 | -12.9756
7.4594 | -34.2563 | -12.9855
7.4655 | -34.2567 | -12.9951
7.4716 | -34.257 | -13.0049
7.4777 | -34.2573 | -13.0146
7.4838 | -34.2577 | -13.0242
7.4899 | -34.2581 | -13.0338
7.496 | -34.2584 | -13.0434
7.5021 | -34.2588 | -13.053
7.5082 | -34.2591 | -13.0627
7.5143 | -34.2595 | -13.0724
7.5204 | -34.2598 | -13.0822
7.5266 | -34.2602 | -13.0918
7.5327 | -34.2606 | -13.1014
7.5388 | -34.2609 | -13.1114
7.5449 | -34.2613 | -13.1212
7.551 | -34.2616 | -13.1308
7.5571 | -34.262 | -13.1406
7.5632 | -34.2624 | -13.1504
7.5693 | -34.2628 | -13.16
7.5754 | -34.2631 | -13.1695
7.5815 | -34.2635 | -13.1792
7.5876 | -34.2638 | -13.1891
7.5937 | -34.2642 | -13.199
7.5998 | -34.2646 | -13.2086
7.6059 | -34.265 | -13.2185
7.612 | -34.2653 | -13.2282
7.6181 | -34.2657 | -13.2378
7.6242 | -34.2661 | -13.2478
7.6303 | -34.2664 | -13.2577
7.6364 | -34.2668 | -13.2675
7.6425 | -34.2671 | -13.2772
7.6486 | -34.2675 | -13.2869
7.6547 | -34.2679 | -13.2965
7.6608 | -34.2682 | -13.3062
7.667 | -34.2686 | -13.316
7.6731 | -34.2689 | -13.3258
7.6792 | -34.2693 | -13.3358
7.6853 | -34.2697 | -13.3456
7.6914 | -34.27 | -13.3554
7.6975 | -34.2704 | -13.3652
7.7036 | -34.2708 | -13.3751
7.7097 | -34.2712 | -13.3849
7.7158 | -34.2715 | -13.3948
7.7219 | -34.2719 | -13.4043
7.728 | -34.2723 | -13.4142
7.7341 | -34.2726 | -13.4241
7.7402 | -34.273 | -13.4339
7.7463 | -34.2734 | -13.4437
7.7524 | -34.2737 | -13.4537
7.7585 | -34.2741 | -13.4635
7.7646 | -34.2745 | -13.4733
7.7707 | -34.2748 | -13.4833
7.7768 | -34.2752 | -13.4932
7.7829 | -34.2755 | -13.5032
7.789 | -34.2759 | -13.5133
7.7951 | -34.2763 | -13.523
7.8012 | -34.2766 | -13.5328
7.8073 | -34.277 | -13.5428
7.8135 | -34.2774 | -13.5526
7.8196 | -34.2778 | -13.5626
7.8257 | -34.2781 | -13.5726
7.8318 | -34.2785 | -13.5825
7.8379 | -34.2789 | -13.5923
7.844 | -34.2792 | -13.602
7.8501 | -34.2796 | -13.612
7.8562 | -34.28 | -13.622
7.8623 | -34.2804 | -13.6317
7.8684 | -34.2807 | -13.6416
7.8745 | -34.2811 | -13.6515
7.8806 | -34.2815 | -13.6615
7.8867 | -34.2818 | -13.6711
7.8928 | -34.2822 | -13.6807
7.8989 | -34.2826 | -13.6906
7.905 | -34.283 | -13.7005
7.9111 | -34.2833 | -13.7104
7.9172 | -34.2837 | -13.7202
7.9233 | -34.284 | -13.7301
7.9294 | -34.2844 | -13.74
7.9355 | -34.2848 | -13.7496
7.9416 | -34.2851 | -13.7595
7.9477 | -34.2855 | -13.7692
7.9539 | -34.2859 | -13.779
7.96 | -34.2862 | -13.7889
7.9661 | -34.2866 | -13.7989
7.9722 | -34.287 | -13.8086
7.9783 | -34.2874 | -13.8186
7.9844 | -34.2878 | -13.8285
7.9905 | -34.2882 | -13.8385
7.9966 | -34.2885 | -13.8485
8.0027 | -34.2889 | -13.8584
8.0088 | -34.2892 | -13.8682
8.0149 | -34.2896 | -13.8781
8.021 | -34.2899 | -13.888
8.0271 | -34.2903 | -13.8981
8.0332 | -34.2907 | -13.9079
8.0393 | -34.2911 | -13.9179
8.0454 | -34.2915 | -13.928
8.0515 | -34.2919 | -13.9379
8.0576 | -34.2922 | -13.948
8.0637 | -34.2926 | -13.9578
8.0698 | -34.2929 | -13.9675
8.0759 | -34.2933 | -13.9773
8.082 | -34.2937 | -13.987
8.0881 | -34.294 | -13.997
8.0942 | -34.2944 | -14.0069
8.1004 | -34.2948 | -14.0169
8.1065 | -34.2952 | -14.0266
8.1126 | -34.2955 | -14.0365
8.1187 | -34.2959 | -14.0465
8.1248 | -34.2962 | -14.0561
8.1309 | -34.2966 | -14.0658
8.137 | -34.297 | -14.0755
8.1431 | -34.2974 | -14.0856
8.1492 | -34.2977 | -14.0955
8.1553 | -34.2981 | -14.1054
8.1614 | -34.2985 | -14.1155
8.1675 | -34.2988 | -14.1253
8.1736 | -34.2992 | -14.1354
8.1797 | -34.2995 | -14.1454
8.1858 | -34.2999 | -14.1553
8.1919 | -34.3003 | -14.1651
8.198 | -34.3006 | -14.1748
8.2041 | -34.3011 | -14.1848
8.2102 | -34.3014 | -14.1947
8.2163 | -34.3018 | -14.2043
8.2224 | -34.3022 | -14.2142
8.2285 | -34.3025 | -14.2238
8.2346 | -34.3029 | -14.2337
8.2408 | -34.3033 | -14.2433
8.2469 | -34.3036 | -14.2533
8.253 | -34.304 | -14.263
8.2591 | -34.3044 | -14.2728
8.2652 | -34.3048 | -14.2827
8.2713 | -34.3052 | -14.2926
8.2774 | -34.3056 | -14.3025
8.2835 | -34.306 | -14.3127
8.2896 | -34.3063 | -14.3224
8.2957 | -34.3067 | -14.3323
8.3018 | -34.3071 | -14.3421
8.3079 | -34.3075 | -14.3518
8.314 | -34.3079 | -14.3616
8.3201 | -34.3082 | -14.3714
8.3262 | -34.3086 | -14.3813
8.3323 | -34.309 | -14.3911
8.3384 | -34.3094 | -14.4011
8.3445 | -34.3097 | -14.4109
8.3506 | -34.3101 | -14.4207
8.3567 | -34.3105 | -14.4307
8.3628 | -34.3109 | -14.4405
8.3689 | -34.3112 | -14.4502
8.375 | -34.3116 | -14.4602
8.3812 | -34.3119 | -14.4699
8.3873 | -34.3123 | -14.4798
8.3934 | -34.3127 | -14.4897
8.3995 | -34.3131 | -14.4995
8.4056 | -34.3134 | -14.5093
8.4117 | -34.3138 | -14.5192
8.4178 | -34.3142 | -14.5291
8.4239 | -34.3145 | -14.5387
8.43 | -34.3149 | -14.5486
8.4361 | -34.3152 | -14.5586
8.4422 | -34.3156 | -14.5683
8.4483 | -34.316 | -14.578
8.4544 | -34.3164 | -14.5879
8.4605 | -34.3168 | -14.5976
8.4666 | -34.3171 | -14.6076
8.4727 | -34.3175 | -14.6175
8.4788 | -34.3178 | -14.6273
8.4849 | -34.3182 | -14.6371
8.491 | -34.3186 | -14.6469
8.4971 | -34.319 | -14.6565
8.5032 | -34.3193 | -14.6663
8.5093 | -34.3197 | -14.676
8.5154 | -34.3201 | -14.6858
8.5215 | -34.3204 | -14.6956
8.5277 | -34.3208 | -14.7053
8.5338 | -34.3212 | -14.7151
8.5399 | -34.3215 | -14.7249
8.546 | -34.3219 | -14.7345
8.5521 | -34.3223 | -14.7443
8.5582 | -34.3226 | -14.754
8.5643 | -34.323 | -14.7636
8.5704 | -34.3234 | -14.7731
8.5765 | -34.3238 | -14.783
8.5826 | -34.3242 | -14.7924
8.5887 | -34.3246 | -14.8022
8.5948 | -34.3249 | -14.812
8.6009 | -34.3253 | -14.8216
8.607 | -34.3257 | -14.8315
8.6131 | -34.3261 | -14.8414
8.6192 | -34.3265 | -14.851
8.6253 | -34.3268 | -14.8606
8.6314 | -34.3272 | -14.8703
8.6375 | -34.3275 | -14.8801
8.6436 | -34.3279 | -14.8899
8.6497 | -34.3283 | -14.8994
8.6558 | -34.3286 | -14.9092
8.6619 | -34.329 | -14.919
8.6681 | -34.3293 | -14.9288
8.6742 | -34.3297 | -14.9385
8.6803 | -34.3301 | -14.9482
8.6864 | -34.3305 | -14.9576
8.6925 | -34.3309 | -14.9671
8.6986 | -34.3312 | -14.9769
8.7047 | -34.3316 | -14.9868
8.7108 | -34.3319 | -14.9965
8.7169 | -34.3323 | -15.0063
8.723 | -34.3327 | -15.0158
8.7291 | -34.3331 | -15.0253
8.7352 | -34.3334 | -15.0353
8.7413 | -34.3338 | -15.0451
8.7474 | -34.3342 | -15.0546
8.7535 | -34.3346 | -15.0643
8.7596 | -34.335 | -15.074
8.7657 | -34.3353 | -15.0838
8.7718 | -34.3357 | -15.0935
8.7779 | -34.3361 | -15.1033
8.784 | -34.3364 | -15.1131
8.7901 | -34.3368 | -15.1227
8.7962 | -34.3372 | -15.1321
8.8023 | -34.3376 | -15.1418
8.8084 | -34.3379 | -15.1514
8.8146 | -34.3383 | -15.161
8.8207 | -34.3387 | -15.1706
8.8268 | -34.3391 | -15.1799
8.8329 | -34.3394 | -15.1896
8.839 | -34.3398 | -15.1993
8.8451 | -34.3402 | -15.2089
8.8512 | -34.3406 | -15.2184
8.8573 | -34.3409 | -15.2281
8.8634 | -34.3413 | -15.2375
8.8695 | -34.3417 | -15.2473
8.8756 | -34.342 | -15.257
8.8817 | -34.3424 | -15.2665
8.8878 | -34.3428 | -15.2763
8.8939 | -34.3432 | -15.2859
8.9 | -34.3436 | -15.2955
8.9061 | -34.3439 | -15.3051
8.9122 | -34.3443 | -15.3147
8.9183 | -34.3447 | -15.3244
8.9244 | -34.3451 | -15.3341
8.9305 | -34.3455 | -15.3435
8.9366 | -34.3459 | -15.3532
8.9427 | -34.3462 | -15.3628
8.9488 | -34.3466 | -15.3724
8.955 | -34.347 | -15.3819
8.9611 | -34.3474 | -15.3916
8.9672 | -34.3478 | -15.4011
8.9733 | -34.3482 | -15.4106
8.9794 | -34.3486 | -15.42
8.9855 | -34.3489 | -15.4297
8.9916 | -34.3493 | -15.4395
8.9977 | -34.3497 | -15.449
9.0038 | -34.3501 | -15.4585
9.0099 | -34.3505 | -15.4681
9.016 | -34.3509 | -15.4777
9.0221 | -34.3512 | -15.4872
9.0282 | -34.3516 | -15.4967
9.0343 | -34.352 | -15.5063
9.0404 | -34.3525 | -15.5159
9.0465 | -34.3528 | -15.5256
9.0526 | -34.3532 | -15.5351
9.0587 | -34.3537 | -15.5447
9.0648 | -34.3541 | -15.5542
9.0709 | -34.3544 | -15.5638
9.077 | -34.3548 | -15.5734
9.0831 | -34.3552 | -15.5831
9.0892 | -34.3556 | -15.5927
9.0953 | -34.356 | -15.6024
9.1015 | -34.3564 | -15.6119
9.1076 | -34.3568 | -15.6213
9.1137 | -34.3572 | -15.6308
9.1198 | -34.3576 | -15.6404
9.1259 | -34.358 | -15.6499
9.132 | -34.3584 | -15.6597
9.1381 | -34.3588 | -15.6691
9.1442 | -34.3592 | -15.6783
9.1503 | -34.3596 | -15.688
9.1564 | -34.36 | -15.6975
9.1625 | -34.3604 | -15.7071
9.1686 | -34.3608 | -15.7165
9.1747 | -34.3612 | -15.7261
9.1808 | -34.3616 | -15.7356
9.1869 | -34.362 | -15.7451
9.193 | -34.3624 | -15.7546
9.1991 | -34.3628 | -15.7642
9.2052 | -34.3632 | -15.7738
9.2113 | -34.3636 | -15.7833
9.2174 | -34.3641 | -15.7928
9.2235 | -34.3645 | -15.8023
9.2296 | -34.3649 | -15.8118
9.2357 | -34.3653 | -15.8214
9.2419 | -34.3657 | -15.8309
9.248 | -34.3661 | -15.8404
9.2541 | -34.3665 | -15.8499
9.2602 | -34.3669 | -15.8596
9.2663 | -34.3673 | -15.8689
9.2724 | -34.3678 | -15.8783
9.2785 | -34.3681 | -15.8878
9.2846 | -34.3686 | -15.8971
9.2907 | -34.369 | -15.9066
9.2968 | -34.3694 | -15.9162
9.3029 | -34.3698 | -15.9259
9.309 | -34.3702 | -15.9354
9.3151 | -34.3706 | -15.9449
9.3212 | -34.371 | -15.9542
9.3273 | -34.3714 | -15.9637
9.3334 | -34.3718 | -15.973
9.3395 | -34.3722 | -15.9824
9.3456 | -34.3726 | -15.9917
9.3517 | -34.373 | -16.0013
9.3578 | -34.3734 | -16.0107
9.3639 | -34.3738 | -16.0201
9.37 | -34.3743 | -16.0296
9.3761 | -34.3747 | -16.0392
9.3822 | -34.3751 | -16.0487
9.3884 | -34.3755 | -16.0579
9.3945 | -34.3759 | -16.0674
9.4006 | -34.3764 | -16.0769
9.4067 | -34.3767 | -16.0865
9.4128 | -34.3772 | -16.096
9.4189 | -34.3776 | -16.1057
9.425 | -34.378 | -16.115
9.4311 | -34.3784 | -16.1245
9.4372 | -34.3789 | -16.134
9.4433 | -34.3793 | -16.1434
9.4494 | -34.3797 | -16.1527
9.4555 | -34.3801 | -16.1623
9.4616 | -34.3805 | -16.1717
9.4677 | -34.381 | -16.1812
9.4738 | -34.3814 | -16.1906
9.4799 | -34.3818 | -16.2001
9.486 | -34.3822 | -16.2096
9.4921 | -34.3826 | -16.219
9.4982 | -34.383 | -16.2283
9.5043 | -34.3835 | -16.2377
9.5104 | -34.3839 | -16.247
9.5165 | -34.3843 | -16.2564
9.5226 | -34.3847 | -16.2659
9.5288 | -34.3851 | -16.2752
9.5349 | -34.3855 | -16.2846
9.541 | -34.386 | -16.2939
9.5471 | -34.3864 | -16.3033
9.5532 | -34.3868 | -16.3128
9.5593 | -34.3872 | -16.3223
9.5654 | -34.3876 | -16.3317
9.5715 | -34.388 | -16.3411
9.5776 | -34.3884 | -16.3507
9.5837 | -34.3888 | -16.3601
9.5898 | -34.3892 | -16.3697
9.5959 | -34.3897 | -16.3791
9.602 | -34.3901 | -16.3884
9.6081 | -34.3905 | -16.3978
9.6142 | -34.3909 | -16.4071
9.6203 | -34.3913 | -16.4166
9.6264 | -34.3917 | -16.4259
9.6325 | -34.3921 | -16.4354
9.6386 | -34.3926 | -16.4446
9.6447 | -34.393 | -16.4541
9.6508 | -34.3934 | -16.4634
9.6569 | -34.3938 | -16.4728
9.663 | -34.3942 | -16.4823
9.6691 | -34.3947 | -16.4917
9.6753 | -34.3951 | -16.501
9.6814 | -34.3955 | -16.5103
9.6875 | -34.3959 | -16.5195
9.6936 | -34.3963 | -16.5288
9.6997 | -34.3968 | -16.5383
9.7058 | -34.3972 | -16.5476
9.7119 | -34.3977 | -16.557
9.718 | -34.3981 | -16.5666
9.7241 | -34.3985 | -16.5758
9.7302 | -34.3989 | -16.5853
9.7363 | -34.3993 | -16.5946
9.7424 | -34.3998 | -16.6039
9.7485 | -34.4002 | -16.6134
9.7546 | -34.4006 | -16.6226
9.7607 | -34.401 | -16.6322
9.7668 | -34.4014 | -16.6417
9.7729 | -34.4019 | -16.651
9.779 | -34.4023 | -16.6602
9.7851 | -34.4027 | -16.6697
9.7912 | -34.4032 | -16.679
9.7973 | -34.4036 | -16.6885
9.8034 | -34.404 | -16.6977
9.8095 | -34.4045 | -16.7073
9.8157 | -34.4049 | -16.7167
9.8218 | -34.4053 | -16.7262
9.8279 | -34.4057 | -16.7355
9.834 | -34.4061 | -16.7449
9.8401 | -34.4066 | -16.7543
9.8462 | -34.407 | -16.7635
9.8523 | -34.4074 | -16.7729
9.8584 | -34.4078 | -16.7822
9.8645 | -34.4082 | -16.7916
9.8706 | -34.4086 | -16.801
9.8767 | -34.4091 | -16.8103
9.8828 | -34.4095 | -16.8195
9.8889 | -34.4099 | -16.8287
9.895 | -34.4103 | -16.8381
9.9011 | -34.4107 | -16.8474
9.9072 | -34.4111 | -16.8568
9.9133 | -34.4116 | -16.8661
9.9194 | -34.412 | -16.8754
9.9255 | -34.4124 | -16.8846
9.9316 | -34.4128 | -16.8939
9.9377 | -34.4132 | -16.9033
9.9438 | -34.4137 | -16.9126
9.9499 | -34.4141 | -16.9222
9.956 | -34.4145 | -16.9315
9.9622 | -34.4149 | -16.941
9.9683 | -34.4153 | -16.9503
9.9744 | -34.4157 | -16.9598
9.9805 | -34.4162 | -16.9691
9.9866 | -34.4166 | -16.9784
9.9927 | -34.417 | -16.9876
9.9988 | -34.4174 | -16.997
10.0049 | -34.4179 | -17.0064
10.011 | -34.4183 | -17.0158
10.0171 | -34.4187 | -17.025
10.0232 | -34.4191 | -17.0343
10.0293 | -34.4196 | -17.0437
10.0354 | -34.42 | -17.053
10.0415 | -34.4205 | -17.0624
10.0476 | -34.4209 | -17.0718
10.0537 | -34.4213 | -17.0813
10.0598 | -34.4217 | -17.0907
10.0659 | -34.4221 | -17.1
10.072 | -34.4225 | -17.1091
10.0781 | -34.423 | -17.1185
10.0842 | -34.4234 | -17.1278
10.0903 | -34.4238 | -17.1371
10.0964 | -34.4242 | -17.1462
10.1026 | -34.4246 | -17.1555
10.1087 | -34.425 | -17.1648
10.1148 | -34.4254 | -17.1742
10.1209 | -34.4259 | -17.1834
10.127 | -34.4263 | -17.1925
10.1331 | -34.4267 | -17.2018
10.1392 | -34.4271 | -17.2109
10.1453 | -34.4276 | -17.2202
10.1514 | -34.428 | -17.2295
10.1575 | -34.4284 | -17.239
10.1636 | -34.4288 | -17.2479
10.1697 | -34.4292 | -17.2574
10.1758 | -34.4297 | -17.2665
10.1819 | -34.4301 | -17.2758
10.188 | -34.4305 | -17.2852
10.1941 | -34.4309 | -17.2944
10.2002 | -34.4313 | -17.3038
10.2063 | -34.4317 | -17.313
10.2124 | -34.4321 | -17.3224
10.2185 | -34.4326 | -17.3318
10.2246 | -34.433 | -17.3409
10.2307 | -34.4334 | -17.3502
10.2368 | -34.4338 | -17.3592
10.2429 | -34.4342 | -17.3685
10.2491 | -34.4347 | -17.3779
10.2552 | -34.4351 | -17.3871
10.2613 | -34.4355 | -17.3964
10.2674 | -34.4359 | -17.4058
10.2735 | -34.4363 | -17.4152
10.2796 | -34.4368 | -17.4244
10.2857 | -34.4372 | -17.4338
10.2918 | -34.4376 | -17.4433
10.2979 | -34.438 | -17.4526
10.304 | -34.4385 | -17.4616
10.3101 | -34.4389 | -17.4709
10.3162 | -34.4393 | -17.4804
10.3223 | -34.4397 | -17.4895
10.3284 | -34.4401 | -17.4987
10.3345 | -34.4405 | -17.5078
10.3406 | -34.4409 | -17.517
10.3467 | -34.4413 | -17.5263
10.3528 | -34.4417 | -17.5354
10.3589 | -34.4422 | -17.5447
10.365 | -34.4426 | -17.554
10.3711 | -34.443 | -17.5633
10.3772 | -34.4434 | -17.5725
10.3833 | -34.4439 | -17.5818
10.3895 | -34.4443 | -17.591
10.3956 | -34.4447 | -17.6003
10.4017 | -34.4451 | -17.6093
10.4078 | -34.4455 | -17.6185
10.4139 | -34.446 | -17.6279
10.42 | -34.4464 | -17.6373
10.4261 | -34.4468 | -17.6465
10.4322 | -34.4472 | -17.6557
10.4383 | -34.4476 | -17.6649
10.4444 | -34.448 | -17.674
10.4505 | -34.4484 | -17.6833
10.4566 | -34.4488 | -17.6924
10.4627 | -34.4493 | -17.7016
10.4688 | -34.4497 | -17.7107
10.4749 | -34.4501 | -17.72
10.481 | -34.4505 | -17.7293
10.4871 | -34.451 | -17.7384
10.4932 | -34.4514 | -17.7475
10.4993 | -34.4518 | -17.7568
10.5054 | -34.4523 | -17.7661
10.5115 | -34.4527 | -17.7753
10.5176 | -34.4531 | -17.7843
10.5237 | -34.4535 | -17.7936
10.5298 | -34.4539 | -17.8027
10.536 | -34.4543 | -17.8121
10.5421 | -34.4548 | -17.8212
10.5482 | -34.4552 | -17.8305
10.5543 | -34.4556 | -17.8397
10.5604 | -34.456 | -17.8488
10.5665 | -34.4564 | -17.8578
10.5726 | -34.4568 | -17.8671
10.5787 | -34.4572 | -17.8763
10.5848 | -34.4576 | -17.8852
10.5909 | -34.4581 | -17.8945
10.597 | -34.4585 | -17.9038
10.6031 | -34.4589 | -17.9132
10.6092 | -34.4593 | -17.9224
10.6153 | -34.4597 | -17.9317
10.6214 | -34.4601 | -17.9408
10.6275 | -34.4606 | -17.95
10.6336 | -34.461 | -17.9592
10.6397 | -34.4613 | -17.9684
10.6458 | -34.4618 | -17.9775
10.6519 | -34.4622 | -17.9868
10.658 | -34.4626 | -17.996
10.6641 | -34.463 | -18.0052
10.6702 | -34.4634 | -18.0144
10.6764 | -34.4638 | -18.0238
10.6825 | -34.4643 | -18.0328
10.6886 | -34.4647 | -18.0418
10.6947 | -34.4651 | -18.0511
10.7008 | -34.4655 | -18.0603
10.7069 | -34.466 | -18.0695
10.713 | -34.4664 | -18.0787
10.7191 | -34.4668 | -18.0878
10.7252 | -34.4671 | -18.0968
10.7313 | -34.4675 | -18.1059
10.7374 | -34.4679 | -18.115
10.7435 | -34.4683 | -18.1243
10.7496 | -34.4687 | -18.1333
10.7557 | -34.4691 | -18.1424
10.7618 | -34.4696 | -18.1513
10.7679 | -34.47 | -18.1605
10.774 | -34.4704 | -18.1697
10.7801 | -34.4708 | -18.1789
10.7862 | -34.4712 | -18.1879
10.7923 | -34.4716 | -18.1971
10.7984 | -34.472 | -18.206
10.8045 | -34.4725 | -18.2149
10.8106 | -34.4729 | -18.2242
10.8168 | -34.4733 | -18.2333
10.8229 | -34.4737 | -18.2425
10.829 | -34.4741 | -18.2515
10.8351 | -34.4745 | -18.2606
10.8412 | -34.475 | -18.2697
10.8473 | -34.4754 | -18.2786
10.8534 | -34.4758 | -18.2877
10.8595 | -34.4762 | -18.2969
10.8656 | -34.4767 | -18.3058
10.8717 | -34.4771 | -18.3149
10.8778 | -34.4775 | -18.324
10.8839 | -34.4779 | -18.3332
10.89 | -34.4784 | -18.3423
10.8961 | -34.4788 | -18.3513
10.9022 | -34.4792 | -18.3604
10.9083 | -34.4796 | -18.3694
10.9144 | -34.48 | -18.3785
10.9205 | -34.4804 | -18.3876
10.9266 | -34.4808 | -18.3966
10.9327 | -34.4813 | -18.4056
10.9388 | -34.4817 | -18.4145
10.9449 | -34.4821 | -18.4235
10.951 | -34.4825 | -18.4324
10.9571 | -34.4829 | -18.4415
10.9633 | -34.4833 | -18.4507
10.9694 | -34.4838 | -18.4601
10.9755 | -34.4842 | -18.4692
10.9816 | -34.4846 | -18.4784
10.9877 | -34.485 | -18.4873
10.9938 | -34.4854 | -18.4966
10.9999 | -34.4859 | -18.5054
11.006 | -34.4863 | -18.5143
11.0121 | -34.4867 | -18.5234
11.0182 | -34.4871 | -18.5325
11.0243 | -34.4876 | -18.5415
11.0304 | -34.488 | -18.5506
11.0365 | -34.4884 | -18.5596
11.0426 | -34.4888 | -18.5685
11.0487 | -34.4893 | -18.5777
11.0548 | -34.4897 | -18.5866
11.0609 | -34.4901 | -18.5957
11.067 | -34.4905 | -18.6049
11.0731 | -34.491 | -18.6139
11.0792 | -34.4914 | -18.6228
11.0853 | -34.4918 | -18.6318
11.0914 | -34.4922 | -18.6409
11.0975 | -34.4926 | -18.65
11.1037 | -34.493 | -18.6591
11.1098 | -34.4934 | -18.6682
11.1159 | -34.4939 | -18.6771
11.122 | -34.4943 | -18.6862
11.1281 | -34.4947 | -18.6952
11.1342 | -34.4951 | -18.7041
11.1403 | -34.4956 | -18.7133
11.1464 | -34.496 | -18.7224
11.1525 | -34.4964 | -18.7314
11.1586 | -34.4968 | -18.7404
11.1647 | -34.4972 | -18.7494
11.1708 | -34.4977 | -18.7584
11.1769 | -34.4981 | -18.7675
11.183 | -34.4985 | -18.7763
11.1891 | -34.4989 | -18.7853
11.1952 | -34.4994 | -18.7943
11.2013 | -34.4998 | -18.8034
11.2074 | -34.5002 | -18.8125
11.2135 | -34.5006 | -18.8216
11.2196 | -34.5011 | -18.8305
11.2257 | -34.5015 | -18.8396
11.2318 | -34.5019 | -18.8484
11.2379 | -34.5024 | -18.8576
11.244 | -34.5028 | -18.8668
11.2502 | -34.5032 | -18.8757
11.2563 | -34.5037 | -18.8846
11.2624 | -34.5041 | -18.8934
11.2685 | -34.5045 | -18.9022
11.2746 | -34.5049 | -18.9113
11.2807 | -34.5054 | -18.9204
11.2868 | -34.5058 | -18.9297
11.2929 | -34.5062 | -18.9383
11.299 | -34.5066 | -18.9475
11.3051 | -34.507 | -18.9568
11.3112 | -34.5075 | -18.9659
11.3173 | -34.5079 | -18.9748
11.3234 | -34.5084 | -18.9835
11.3295 | -34.5088 | -18.9926
11.3356 | -34.5092 | -19.0016
11.3417 | -34.5097 | -19.0104
11.3478 | -34.5101 | -19.0196
11.3539 | -34.5105 | -19.0288
11.36 | -34.5109 | -19.0375
11.3661 | -34.5114 | -19.0463
11.3722 | -34.5118 | -19.0551
11.3783 | -34.5122 | -19.0641
11.3844 | -34.5127 | -19.0731
11.3906 | -34.5131 | -19.0821
11.3967 | -34.5136 | -19.0909
11.4028 | -34.514 | -19.1001
11.4089 | -34.5144 | -19.109
11.415 | -34.5148 | -19.1178
11.4211 | -34.5152 | -19.1267
11.4272 | -34.5157 | -19.1357
11.4333 | -34.5161 | -19.1446
11.4394 | -34.5165 | -19.1534
11.4455 | -34.5169 | -19.1622
11.4516 | -34.5173 | -19.1712
11.4577 | -34.5178 | -19.1801
11.4638 | -34.5183 | -19.189
11.4699 | -34.5187 | -19.1981
11.476 | -34.5191 | -19.2072
11.4821 | -34.5196 | -19.2163
11.4882 | -34.52 | -19.2254
11.4943 | -34.5204 | -19.2345
11.5004 | -34.5209 | -19.2432
11.5065 | -34.5213 | -19.2521
11.5126 | -34.5218 | -19.2609
11.5187 | -34.5222 | -19.2699
11.5248 | -34.5226 | -19.279
11.5309 | -34.5231 | -19.2879
11.5371 | -34.5235 | -19.2966
11.5432 | -34.5239 | -19.3056
11.5493 | -34.5244 | -19.3148
11.5554 | -34.5248 | -19.3238
11.5615 | -34.5252 | -19.3328
11.5676 | -34.5257 | -19.3418
11.5737 | -34.5261 | -19.3507
11.5798 | -34.5266 | -19.3595
11.5859 | -34.527 | -19.3684
11.592 | -34.5274 | -19.3773
11.5981 | -34.5279 | -19.3862
11.6042 | -34.5283 | -19.3952
11.6103 | -34.5287 | -19.404
11.6164 | -34.5292 | -19.4129
11.6225 | -34.5296 | -19.4216
11.6286 | -34.53 | -19.4305
11.6347 | -34.5304 | -19.4395
11.6408 | -34.5309 | -19.4483
11.6469 | -34.5313 | -19.4572
11.653 | -34.5318 | -19.466
11.6591 | -34.5322 | -19.475
11.6652 | -34.5326 | -19.4838
11.6713 | -34.5331 | -19.4926
11.6775 | -34.5335 | -19.5015
11.6836 | -34.534 | -19.5104
11.6897 | -34.5344 | -19.5193
11.6958 | -34.5349 | -19.5283
11.7019 | -34.5353 | -19.5373
11.708 | -34.5357 | -19.5461
11.7141 | -34.5362 | -19.5548
11.7202 | -34.5367 | -19.5637
11.7263 | -34.5371 | -19.5726
11.7324 | -34.5375 | -19.5816
11.7385 | -34.5379 | -19.5906
11.7446 | -34.5384 | -19.5994
11.7507 | -34.5388 | -19.6081
11.7568 | -34.5393 | -19.617
11.7629 | -34.5397 | -19.6259
11.769 | -34.5401 | -19.6346
11.7751 | -34.5405 | -19.6435
11.7812 | -34.541 | -19.6523
11.7873 | -34.5415 | -19.6611
11.7934 | -34.5419 | -19.67
11.7995 | -34.5423 | -19.679
11.8056 | -34.5428 | -19.6877
11.8117 | -34.5432 | -19.6966
11.8178 | -34.5437 | -19.7052
11.824 | -34.5442 | -19.7141
11.8301 | -34.5446 | -19.7229
11.8362 | -34.545 | -19.7316
11.8423 | -34.5455 | -19.7406
11.8484 | -34.5459 | -19.7493
11.8545 | -34.5464 | -19.7581
11.8606 | -34.5468 | -19.7669
11.8667 | -34.5473 | -19.7758
11.8728 | -34.5477 | -19.7844
11.8789 | -34.5482 | -19.7932
11.885 | -34.5486 | -19.802
11.8911 | -34.5491 | -19.8109
11.8972 | -34.5495 | -19.8196
11.9033 | -34.5499 | -19.8284
11.9094 | -34.5504 | -19.8371
11.9155 | -34.5508 | -19.8459
11.9216 | -34.5513 | -19.8545
11.9277 | -34.5518 | -19.8632
11.9338 | -34.5522 | -19.8721
11.9399 | -34.5526 | -19.8809
11.946 | -34.5531 | -19.8897
11.9521 | -34.5535 | -19.8983
11.9582 | -34.554 | -19.9069
11.9644 | -34.5545 | -19.9157
11.9705 | -34.5549 | -19.9244
11.9766 | -34.5554 | -19.933
11.9827 | -34.5558 | -19.9417
11.9888 | -34.5563 | -19.9504
11.9949 | -34.5568 | -19.9588
12.001 | -34.5572 | -19.9676
12.0071 | -34.5577 | -19.9764
12.0132 | -34.5581 | -19.9851
12.0193 | -34.5585 | -19.9939
12.0254 | -34.559 | -20.0027
12.0315 | -34.5594 | -20.0112
12.0376 | -34.5599 | -20.02
12.0437 | -34.5603 | -20.0288
12.0498 | -34.5607 | -20.0376
12.0559 | -34.5612 | -20.0462
12.062 | -34.5617 | -20.0547
12.0681 | -34.5621 | -20.0634
12.0742 | -34.5626 | -20.0721
12.0803 | -34.563 | -20.0811
12.0864 | -34.5635 | -20.0898
12.0925 | -34.5639 | -20.0985
12.0986 | -34.5644 | -20.1071
12.1047 | -34.5648 | -20.1159
12.1109 | -34.5653 | -20.1246
12.117 | -34.5658 | -20.1332
12.1231 | -34.5662 | -20.1419
12.1292 | -34.5667 | -20.1506
12.1353 | -34.5671 | -20.1592
12.1414 | -34.5676 | -20.168
12.1475 | -34.568 | -20.1767
12.1536 | -34.5685 | -20.1855
12.1597 | -34.5689 | -20.1942
12.1658 | -34.5694 | -20.2029
12.1719 | -34.5698 | -20.2114
12.178 | -34.5703 | -20.2201
12.1841 | -34.5707 | -20.2289
12.1902 | -34.5712 | -20.2377
12.1963 | -34.5717 | -20.2465
12.2024 | -34.5722 | -20.2549
12.2085 | -34.5726 | -20.2636
12.2146 | -34.5731 | -20.2723
12.2207 | -34.5735 | -20.2809
12.2268 | -34.574 | -20.2897
12.2329 | -34.5745 | -20.2983
12.239 | -34.575 | -20.3067
12.2451 | -34.5754 | -20.3153
12.2513 | -34.5758 | -20.3241
12.2574 | -34.5763 | -20.3327
12.2635 | -34.5768 | -20.3414
12.2696 | -34.5772 | -20.35
12.2757 | -34.5777 | -20.3584
12.2818 | -34.5782 | -20.367
12.2879 | -34.5786 | -20.3757
12.294 | -34.5791 | -20.3841
12.3001 | -34.5796 | -20.3928
12.3062 | -34.58 | -20.4015
12.3123 | -34.5805 | -20.4098
12.3184 | -34.581 | -20.4184
12.3245 | -34.5814 | -20.4269
12.3306 | -34.5819 | -20.4355
12.3367 | -34.5824 | -20.4441
12.3428 | -34.5828 | -20.4526
12.3489 | -34.5833 | -20.4611
12.355 | -34.5838 | -20.4699
12.3611 | -34.5842 | -20.4787
12.3672 | -34.5847 | -20.4873
12.3733 | -34.5852 | -20.4959
12.3794 | -34.5856 | -20.5042
12.3855 | -34.5861 | -20.5127
12.3916 | -34.5865 | -20.5214
12.3978 | -34.587 | -20.5301
12.4039 | -34.5875 | -20.5389
12.41 | -34.588 | -20.5476
12.4161 | -34.5885 | -20.5558
12.4222 | -34.589 | -20.5643
12.4283 | -34.5894 | -20.5728
12.4344 | -34.5899 | -20.5816
12.4405 | -34.5903 | -20.5902
12.4466 | -34.5908 | -20.5987
12.4527 | -34.5913 | -20.6073
12.4588 | -34.5917 | -20.6159
12.4649 | -34.5922 | -20.6245
12.471 | -34.5927 | -20.6332
12.4771 | -34.5931 | -20.6418
12.4832 | -34.5936 | -20.6503
12.4893 | -34.594 | -20.6589
12.4954 | -34.5945 | -20.6674
12.5015 | -34.5949 | -20.6759
12.5076 | -34.5955 | -20.6846
12.5137 | -34.5959 | -20.6932
12.5198 | -34.5964 | -20.7017
12.5259 | -34.5968 | -20.7105
12.532 | -34.5973 | -20.719
12.5382 | -34.5978 | -20.7276
12.5443 | -34.5982 | -20.7362
12.5504 | -34.5987 | -20.7448
12.5565 | -34.5992 | -20.7532
12.5626 | -34.5997 | -20.7619
12.5687 | -34.6001 | -20.7706
12.5748 | -34.6006 | -20.7791
12.5809 | -34.601 | -20.7877
12.587 | -34.6015 | -20.7964
12.5931 | -34.602 | -20.8048
12.5992 | -34.6025 | -20.8133
12.6053 | -34.603 | -20.8219
12.6114 | -34.6034 | -20.8305
12.6175 | -34.6039 | -20.8391
12.6236 | -34.6044 | -20.8478
12.6297 | -34.6049 | -20.8563
12.6358 | -34.6053 | -20.8646
12.6419 | -34.6058 | -20.8732
12.648 | -34.6063 | -20.8818
12.6541 | -34.6067 | -20.8903
12.6602 | -34.6072 | -20.8984
12.6663 | -34.6076 | -20.9068
12.6724 | -34.6081 | -20.9153
12.6785 | -34.6086 | -20.9238
12.6847 | -34.6091 | -20.9322
12.6908 | -34.6095 | -20.9408
12.6969 | -34.61 | -20.9491
12.703 | -34.6105 | -20.9577
12.7091 | -34.611 | -20.9662
12.7152 | -34.6114 | -20.9746
12.7213 | -34.6119 | -20.9833
12.7274 | -34.6124 | -20.9918
12.7335 | -34.6129 | -21.0001
12.7396 | -34.6133 | -21.0085
12.7457 | -34.6138 | -21.0171
12.7518 | -34.6143 | -21.0255
12.7579 | -34.6147 | -21.0341
12.764 | -34.6152 | -21.0428
12.7701 | -34.6157 | -21.0511
12.7762 | -34.6161 | -21.0595
12.7823 | -34.6166 | -21.068
12.7884 | -34.6171 | -21.0766
12.7945 | -34.6176 | -21.0852
12.8006 | -34.6181 | -21.0936
12.8067 | -34.6185 | -21.102
12.8128 | -34.619 | -21.1104
12.8189 | -34.6195 | -21.1189
12.8251 | -34.6199 | -21.1274
12.8312 | -34.6204 | -21.136
12.8373 | -34.6209 | -21.1444
12.8434 | -34.6214 | -21.153
12.8495 | -34.6218 | -21.1613
12.8556 | -34.6223 | -21.1697
12.8617 | -34.6227 | -21.1782
12.8678 | -34.6232 | -21.1868
12.8739 | -34.6237 | -21.1951
12.88 | -34.6242 | -21.2036
12.8861 | -34.6246 | -21.2119
12.8922 | -34.6251 | -21.2205
12.8983 | -34.6255 | -21.2289
12.9044 | -34.626 | -21.2372
12.9105 | -34.6265 | -21.2457
12.9166 | -34.6269 | -21.2541
12.9227 | -34.6274 | -21.2626
12.9288 | -34.6279 | -21.2711
12.9349 | -34.6283 | -21.2796
12.941 | -34.6288 | -21.2881
12.9471 | -34.6293 | -21.2966
12.9532 | -34.6298 | -21.3051
12.9593 | -34.6303 | -21.3138
12.9654 | -34.6307 | -21.3222
12.9716 | -34.6312 | -21.3307
12.9777 | -34.6317 | -21.3393
12.9838 | -34.6322 | -21.3476
12.9899 | -34.6327 | -21.356
12.996 | -34.6331 | -21.3645
13.0021 | -34.6336 | -21.3729
13.0082 | -34.6341 | -21.3812
13.0143 | -34.6346 | -21.3895
13.0204 | -34.6351 | -21.3978
13.0265 | -34.6356 | -21.4062
13.0326 | -34.636 | -21.4146
13.0387 | -34.6365 | -21.4232
13.0448 | -34.637 | -21.4317
13.0509 | -34.6375 | -21.4401
13.057 | -34.638 | -21.4484
13.0631 | -34.6385 | -21.4568
13.0692 | -34.639 | -21.4651
13.0753 | -34.6395 | -21.4737
13.0814 | -34.6399 | -21.4822
13.0875 | -34.6404 | -21.4906
13.0936 | -34.6409 | -21.4992
13.0997 | -34.6413 | -21.5074
13.1058 | -34.6418 | -21.5158
13.112 | -34.6423 | -21.5242
13.1181 | -34.6428 | -21.5324
13.1242 | -34.6433 | -21.5409
13.1303 | -34.6438 | -21.5494
13.1364 | -34.6442 | -21.5579
13.1425 | -34.6447 | -21.5662
13.1486 | -34.6451 | -21.5747
13.1547 | -34.6456 | -21.5832
13.1608 | -34.6461 | -21.5916
13.1669 | -34.6466 | -21.5998
13.173 | -34.6471 | -21.6083
13.1791 | -34.6476 | -21.6166
13.1852 | -34.648 | -21.6251
13.1913 | -34.6485 | -21.6335
13.1974 | -34.649 | -21.6421
13.2035 | -34.6494 | -21.6505
13.2096 | -34.6499 | -21.6587
13.2157 | -34.6504 | -21.6671
13.2218 | -34.6508 | -21.6753
13.2279 | -34.6513 | -21.6834
13.234 | -34.6518 | -21.6918
13.2401 | -34.6523 | -21.7001
13.2462 | -34.6527 | -21.7084
13.2524 | -34.6532 | -21.7168
13.2585 | -34.6537 | -21.7251
13.2646 | -34.6542 | -21.7336
13.2707 | -34.6547 | -21.7421
13.2768 | -34.6552 | -21.7505
13.2829 | -34.6556 | -21.7591
13.289 | -34.6561 | -21.7673
13.2951 | -34.6566 | -21.7755
13.3012 | -34.6571 | -21.7839
13.3073 | -34.6576 | -21.7925
13.3134 | -34.658 | -21.8008
13.3195 | -34.6585 | -21.8091
13.3256 | -34.6589 | -21.8175
13.3317 | -34.6594 | -21.8258
13.3378 | -34.6599 | -21.8342
13.3439 | -34.6604 | -21.8427
13.35 | -34.6608 | -21.8512
13.3561 | -34.6613 | -21.8594
13.3622 | -34.6618 | -21.8678
13.3683 | -34.6622 | -21.8759
13.3744 | -34.6627 | -21.8842
13.3805 | -34.6632 | -21.8926
13.3866 | -34.6636 | -21.9008
13.3927 | -34.6641 | -21.9092
13.3989 | -34.6645 | -21.9175
13.405 | -34.665 | -21.9258
13.4111 | -34.6655 | -21.934
13.4172 | -34.6659 | -21.9424
13.4233 | -34.6664 | -21.9509
13.4294 | -34.6669 | -21.9595
13.4355 | -34.6673 | -21.9678
13.4416 | -34.6678 | -21.976
13.4477 | -34.6683 | -21.9844
13.4538 | -34.6688 | -21.9927
13.4599 | -34.6693 | -22.0011
13.466 | -34.6697 | -22.0095
13.4721 | -34.6702 | -22.0177
13.4782 | -34.6707 | -22.026
13.4843 | -34.6711 | -22.0346
13.4904 | -34.6716 | -22.0427
13.4965 | -34.6721 | -22.051
13.5026 | -34.6725 | -22.0594
13.5087 | -34.673 | -22.0676
13.5148 | -34.6735 | -22.0761
13.5209 | -34.674 | -22.0845
13.527 | -34.6745 | -22.0928
13.5331 | -34.675 | -22.1012
13.5393 | -34.6754 | -22.1094
13.5454 | -34.6759 | -22.1175
13.5515 | -34.6764 | -22.126
13.5576 | -34.6768 | -22.1344
13.5637 | -34.6773 | -22.143
13.5698 | -34.6778 | -22.1514
13.5759 | -34.6783 | -22.1594
13.582 | -34.6787 | -22.1677
13.5881 | -34.6792 | -22.176
13.5942 | -34.6797 | -22.1845
13.6003 | -34.6802 | -22.1927
13.6064 | -34.6806 | -22.2012
13.6125 | -34.6811 | -22.2096
13.6186 | -34.6816 | -22.2179
13.6247 | -34.6821 | -22.2263
13.6308 | -34.6826 | -22.2346
13.6369 | -34.683 | -22.243
13.643 | -34.6835 | -22.2515
13.6491 | -34.6839 | -22.2596
13.6552 | -34.6845 | -22.268
13.6613 | -34.6849 | -22.2764
13.6674 | -34.6854 | -22.2846
13.6735 | -34.6859 | -22.2929
13.6796 | -34.6863 | -22.3012
13.6858 | -34.6868 | -22.3096
13.6919 | -34.6873 | -22.318
13.698 | -34.6878 | -22.3263
13.7041 | -34.6883 | -22.3347
13.7102 | -34.6887 | -22.343
13.7163 | -34.6893 | -22.3513
13.7224 | -34.6897 | -22.3596
13.7285 | -34.6902 | -22.3679
13.7346 | -34.6907 | -22.3764
13.7407 | -34.6912 | -22.3848
13.7468 | -34.6916 | -22.393
13.7529 | -34.6921 | -22.4012
13.759 | -34.6926 | -22.4096
13.7651 | -34.693 | -22.418
13.7712 | -34.6935 | -22.4264
13.7773 | -34.694 | -22.4348
13.7834 | -34.6944 | -22.443
13.7895 | -34.6949 | -22.4512
13.7956 | -34.6954 | -22.4596
13.8017 | -34.6958 | -22.4679
13.8078 | -34.6963 | -22.4763
13.8139 | -34.6968 | -22.4845
13.82 | -34.6973 | -22.4927
13.8262 | -34.6977 | -22.501
13.8323 | -34.6982 | -22.5094
13.8384 | -34.6987 | -22.5178
13.8445 | -34.6992 | -22.5261
13.8506 | -34.6996 | -22.5344
13.8567 | -34.7001 | -22.5427
13.8628 | -34.7005 | -22.5511
13.8689 | -34.701 | -22.5595
13.875 | -34.7015 | -22.5677
13.8811 | -34.7019 | -22.5761
13.8872 | -34.7024 | -22.5843
13.8933 | -34.7029 | -22.5926
13.8994 | -34.7033 | -22.6009
13.9055 | -34.7038 | -22.6093
13.9116 | -34.7043 | -22.6176
13.9177 | -34.7047 | -22.626
13.9238 | -34.7052 | -22.6343
13.9299 | -34.7057 | -22.6425
13.936 | -34.7062 | -22.6509
13.9421 | -34.7066 | -22.6591
13.9482 | -34.7071 | -22.6674
13.9543 | -34.7075 | -22.6757
13.9604 | -34.708 | -22.6839
13.9665 | -34.7085 | -22.6922
13.9727 | -34.7089 | -22.7006
13.9788 | -34.7094 | -22.7089
13.9849 | -34.7099 | -22.7172
13.991 | -34.7104 | -22.7256
13.9971 | -34.7108 | -22.7337
14.0032 | -34.7112 | -22.742
14.0093 | -34.7117 | -22.7503
14.0154 | -34.7122 | -22.7584
14.0215 | -34.7127 | -22.7666
14.0276 | -34.7132 | -22.7749
14.0337 | -34.7136 | -22.7832
14.0398 | -34.7141 | -22.7913
14.0459 | -34.7146 | -22.7996
14.052 | -34.7151 | -22.808
14.0581 | -34.7155 | -22.8163
14.0642 | -34.716 | -22.8247
14.0703 | -34.7165 | -22.8329
14.0764 | -34.7169 | -22.841
14.0825 | -34.7174 | -22.8494
14.0886 | -34.7179 | -22.8577
14.0947 | -34.7184 | -22.866
14.1008 | -34.7188 | -22.8744
14.1069 | -34.7193 | -22.8826
14.1131 | -34.7197 | -22.891
14.1192 | -34.7202 | -22.8993
14.1253 | -34.7207 | -22.9073
14.1314 | -34.7211 | -22.9156
14.1375 | -34.7216 | -22.9238
14.1436 | -34.7221 | -22.932
14.1497 | -34.7226 | -22.9403
14.1558 | -34.7231 | -22.9487
14.1619 | -34.7235 | -22.9568
14.168 | -34.724 | -22.965
14.1741 | -34.7245 | -22.9732
14.1802 | -34.7249 | -22.9814
14.1863 | -34.7254 | -22.9896
14.1924 | -34.7259 | -22.9979
14.1985 | -34.7264 | -23.006
14.2046 | -34.7268 | -23.0143
14.2107 | -34.7273 | -23.0224
14.2168 | -34.7278 | -23.0307
14.2229 | -34.7282 | -23.0389
14.229 | -34.7287 | -23.0472
14.2351 | -34.7291 | -23.0555
14.2412 | -34.7296 | -23.0638
14.2473 | -34.7301 | -23.0721
14.2534 | -34.7306 | -23.0802
14.2596 | -34.731 | -23.0883
14.2657 | -34.7315 | -23.0966
14.2718 | -34.7319 | -23.1048
14.2779 | -34.7324 | -23.1131
14.284 | -34.7329 | -23.1212
14.2901 | -34.7334 | -23.1295
14.2962 | -34.7338 | -23.1377
14.3023 | -34.7343 | -23.146
14.3084 | -34.7348 | -23.1543
14.3145 | -34.7353 | -23.1625
14.3206 | -34.7357 | -23.171
14.3267 | -34.7362 | -23.1791
14.3328 | -34.7367 | -23.1873
14.3389 | -34.7371 | -23.1955
14.345 | -34.7376 | -23.2038
14.3511 | -34.738 | -23.2119
14.3572 | -34.7385 | -23.2203
14.3633 | -34.739 | -23.2287
14.3694 | -34.7395 | -23.237
14.3755 | -34.7399 | -23.2453
14.3816 | -34.7404 | -23.2537
14.3877 | -34.7409 | -23.2618
14.3938 | -34.7413 | -23.2702
14.4 | -34.7418 | -23.2784
14.4061 | -34.7423 | -23.2866
14.4122 | -34.7428 | -23.295
14.4183 | -34.7432 | -23.303
14.4244 | -34.7437 | -23.3113
14.4305 | -34.7442 | -23.3195
14.4366 | -34.7447 | -23.3278
14.4427 | -34.7451 | -23.3362
14.4488 | -34.7456 | -23.3445
14.4549 | -34.7461 | -23.3527
14.461 | -34.7466 | -23.3607
14.4671 | -34.747 | -23.369
14.4732 | -34.7475 | -23.3774
14.4793 | -34.7479 | -23.3856
14.4854 | -34.7484 | -23.3939
14.4915 | -34.7489 | -23.4018
14.4976 | -34.7494 | -23.4099
14.5037 | -34.7499 | -23.418
14.5098 | -34.7503 | -23.4263
14.5159 | -34.7508 | -23.4346
14.522 | -34.7512 | -23.4427
14.5281 | -34.7517 | -23.4507
14.5342 | -34.7522 | -23.4591
14.5403 | -34.7527 | -23.4672
14.5465 | -34.7531 | -23.4755
14.5526 | -34.7536 | -23.4839
14.5587 | -34.7541 | -23.4922
14.5648 | -34.7546 | -23.5002
14.5709 | -34.7551 | -23.5082
14.577 | -34.7555 | -23.5165
14.5831 | -34.756 | -23.5248
14.5892 | -34.7565 | -23.5329
14.5953 | -34.7569 | -23.5411
14.6014 | -34.7574 | -23.5493
14.6075 | -34.7579 | -23.5574
14.6136 | -34.7583 | -23.5657
14.6197 | -34.7588 | -23.574
14.6258 | -34.7592 | -23.5821
14.6319 | -34.7597 | -23.5902
14.638 | -34.7602 | -23.5983
14.6441 | -34.7606 | -23.6066
14.6502 | -34.7611 | -23.6148
14.6563 | -34.7616 | -23.623
14.6624 | -34.7621 | -23.6312
14.6685 | -34.7625 | -23.6391
14.6746 | -34.763 | -23.6473
14.6807 | -34.7635 | -23.6554
14.6869 | -34.7639 | -23.6636
14.693 | -34.7644 | -23.6717
14.6991 | -34.7649 | -23.6799
14.7052 | -34.7654 | -23.6878
14.7113 | -34.7659 | -23.696
14.7174 | -34.7663 | -23.7042
14.7235 | -34.7668 | -23.7123
14.7296 | -34.7673 | -23.7205
14.7357 | -34.7678 | -23.7286
14.7418 | -34.7682 | -23.7367
14.7479 | -34.7687 | -23.7448
14.754 | -34.7692 | -23.753
14.7601 | -34.7696 | -23.7611
14.7662 | -34.7701 | -23.7692
14.7723 | -34.7706 | -23.7772
14.7784 | -34.771 | -23.7854
14.7845 | -34.7715 | -23.7936
14.7906 | -34.772 | -23.8018
14.7967 | -34.7724 | -23.8099
14.8028 | -34.7729 | -23.8179
14.8089 | -34.7734 | -23.8259
14.815 | -34.7739 | -23.834
14.8211 | -34.7744 | -23.8422
14.8272 | -34.7748 | -23.8504
14.8334 | -34.7753 | -23.8585
14.8395 | -34.7758 | -23.8667
14.8456 | -34.7763 | -23.8748
14.8517 | -34.7767 | -23.883
14.8578 | -34.7772 | -23.8912
14.8639 | -34.7777 | -23.8993
14.87 | -34.7782 | -23.9074
14.8761 | -34.7786 | -23.9155
14.8822 | -34.7791 | -23.9236
14.8883 | -34.7796 | -23.9317
14.8944 | -34.78 | -23.9397
14.9005 | -34.7805 | -23.9478
14.9066 | -34.781 | -23.956
14.9127 | -34.7815 | -23.9639
14.9188 | -34.7819 | -23.972
14.9249 | -34.7824 | -23.9801
14.931 | -34.7828 | -23.9883
14.9371 | -34.7833 | -23.9964
14.9432 | -34.7838 | -24.0045
14.9493 | -34.7843 | -24.0125
14.9554 | -34.7848 | -24.0207
14.9615 | -34.7853 | -24.0289
14.9676 | -34.7857 | -24.0372
14.9738 | -34.7862 | -24.0451
14.9799 | -34.7867 | -24.0532
14.986 | -34.7871 | -24.0612
14.9921 | -34.7876 | -24.0692
14.9982 | -34.7881 | -24.0773
15.0043 | -34.7885 | -24.0853
15.0104 | -34.789 | -24.0935
15.0165 | -34.7895 | -24.1017
15.0226 | -34.79 | -24.1096
15.0287 | -34.7905 | -24.1177
15.0348 | -34.791 | -24.1258
15.0409 | -34.7914 | -24.1338
15.047 | -34.7919 | -24.142
15.0531 | -34.7924 | -24.1499
15.0592 | -34.7929 | -24.158
15.0653 | -34.7934 | -24.1661
15.0714 | -34.7939 | -24.1741
15.0775 | -34.7943 | -24.182
15.0836 | -34.7948 | -24.1901
15.0897 | -34.7953 | -24.198
15.0958 | -34.7958 | -24.206
15.1019 | -34.7963 | -24.2142
15.108 | -34.7968 | -24.2222
15.1141 | -34.7973 | -24.2303
15.1203 | -34.7978 | -24.2383
15.1264 | -34.7982 | -24.2463
15.1325 | -34.7987 | -24.2543
15.1386 | -34.7992 | -24.2623
15.1447 | -34.7997 | -24.2704
15.1508 | -34.8002 | -24.2784
15.1569 | -34.8007 | -24.2865
15.163 | -34.8012 | -24.2943
15.1691 | -34.8017 | -24.3024
15.1752 | -34.8022 | -24.3102
15.1813 | -34.8026 | -24.3182
15.1874 | -34.8031 | -24.3263
15.1935 | -34.8036 | -24.3341
15.1996 | -34.8041 | -24.3421
15.2057 | -34.8046 | -24.35
15.2118 | -34.8051 | -24.3581
15.2179 | -34.8056 | -24.3662
15.224 | -34.806 | -24.3741
15.2301 | -34.8065 | -24.3818
15.2362 | -34.807 | -24.39
15.2423 | -34.8075 | -24.3981
15.2484 | -34.808 | -24.4061
15.2545 | -34.8085 | -24.4142
15.2607 | -34.809 | -24.4222
15.2668 | -34.8095 | -24.43
15.2729 | -34.81 | -24.4378
15.279 | -34.8105 | -24.446
15.2851 | -34.811 | -24.454
15.2912 | -34.8115 | -24.462
15.2973 | -34.812 | -24.4701
15.3034 | -34.8125 | -24.4779
15.3095 | -34.813 | -24.4858
15.3156 | -34.8134 | -24.4938
15.3217 | -34.8139 | -24.5017
15.3278 | -34.8144 | -24.5099
15.3339 | -34.8149 | -24.5175
15.34 | -34.8154 | -24.5254
15.3461 | -34.8159 | -24.5334
15.3522 | -34.8164 | -24.5414
15.3583 | -34.8169 | -24.5495
15.3644 | -34.8173 | -24.5574
15.3705 | -34.8178 | -24.5649
15.3766 | -34.8183 | -24.5729
15.3827 | -34.8188 | -24.5807
15.3888 | -34.8193 | -24.5886
15.3949 | -34.8198 | -24.5966
15.401 | -34.8203 | -24.6046
15.4072 | -34.8208 | -24.6124
15.4133 | -34.8212 | -24.6203
15.4194 | -34.8217 | -24.6283
15.4255 | -34.8222 | -24.6362
15.4316 | -34.8227 | -24.6442
15.4377 | -34.8232 | -24.6522
15.4438 | -34.8237 | -24.6599
15.4499 | -34.8242 | -24.6678
15.456 | -34.8247 | -24.6757
15.4621 | -34.8252 | -24.6835
15.4682 | -34.8256 | -24.6915
15.4743 | -34.8262 | -24.699
15.4804 | -34.8266 | -24.707
15.4865 | -34.8271 | -24.715
15.4926 | -34.8276 | -24.723
15.4987 | -34.8281 | -24.7309
15.5048 | -34.8286 | -24.7389
15.5109 | -34.8291 | -24.7465
15.517 | -34.8296 | -24.7545
15.5231 | -34.8301 | -24.7624
15.5292 | -34.8306 | -24.7702
15.5353 | -34.8311 | -24.7782
15.5414 | -34.8315 | -24.7861
15.5476 | -34.8321 | -24.7937
15.5537 | -34.8326 | -24.8016
15.5598 | -34.8331 | -24.8096
15.5659 | -34.8336 | -24.8175
15.572 | -34.8341 | -24.8254
15.5781 | -34.8346 | -24.8331
15.5842 | -34.8351 | -24.8407
15.5903 | -34.8356 | -24.8484
15.5964 | -34.8361 | -24.8562
15.6025 | -34.8366 | -24.864
15.6086 | -34.8371 | -24.8718
15.6147 | -34.8376 | -24.8797
15.6208 | -34.8381 | -24.8873
15.6269 | -34.8386 | -24.8951
15.633 | -34.8391 | -24.9031
15.6391 | -34.8396 | -24.911
15.6452 | -34.84 | -24.9189
15.6513 | -34.8406 | -24.9264
15.6574 | -34.8411 | -24.9344
15.6635 | -34.8416 | -24.9423
15.6696 | -34.8421 | -24.9502
15.6757 | -34.8426 | -24.958
15.6818 | -34.843 | -24.9658
15.688 | -34.8435 | -24.9735
15.6941 | -34.844 | -24.9812
15.7002 | -34.8445 | -24.989
15.7063 | -34.845 | -24.9969
15.7124 | -34.8455 | -25.0048
15.7185 | -34.846 | -25.0126
15.7246 | -34.8465 | -25.0202
15.7307 | -34.847 | -25.0278
15.7368 | -34.8475 | -25.0354
15.7429 | -34.8481 | -25.0432
15.749 | -34.8486 | -25.051
15.7551 | -34.8491 | -25.0589
15.7612 | -34.8496 | -25.0665
15.7673 | -34.8501 | -25.0743
15.7734 | -34.8506 | -25.0819
15.7795 | -34.8511 | -25.0897
15.7856 | -34.8516 | -25.0973
15.7917 | -34.8521 | -25.1049
15.7978 | -34.8526 | -25.1128
15.8039 | -34.8531 | -25.1206
15.81 | -34.8536 | -25.1284
15.8161 | -34.8541 | -25.1363
15.8222 | -34.8546 | -25.144
15.8283 | -34.8551 | -25.1515
15.8345 | -34.8556 | -25.1592
15.8406 | -34.8561 | -25.1669
15.8467 | -34.8566 | -25.1746
15.8528 | -34.8571 | -25.1824
15.8589 | -34.8576 | -25.1902
15.865 | -34.8581 | -25.1976
15.8711 | -34.8586 | -25.2053
15.8772 | -34.8591 | -25.2131
15.8833 | -34.8596 | -25.2208
15.8894 | -34.8601 | -25.2286
15.8955 | -34.8606 | -25.2363
15.9016 | -34.8611 | -25.2437
15.9077 | -34.8616 | -25.2514
15.9138 | -34.8621 | -25.2593
15.9199 | -34.8626 | -25.2669
15.926 | -34.8631 | -25.2745
15.9321 | -34.8636 | -25.282
15.9382 | -34.8641 | -25.2898
15.9443 | -34.8646 | -25.2975
15.9504 | -34.865 | -25.3053
15.9565 | -34.8655 | -25.3131
15.9626 | -34.866 | -25.3208
15.9687 | -34.8665 | -25.3282
15.9749 | -34.8671 | -25.3356
15.981 | -34.8676 | -25.3433
15.9871 | -34.8681 | -25.351
15.9932 | -34.8685 | -25.3587
15.9993 | -34.869 | -25.3665
16.0054 | -34.8696 | -25.374
16.0115 | -34.8701 | -25.3818
16.0176 | -34.8706 | -25.3896
16.0237 | -34.8711 | -25.3972
16.0298 | -34.8716 | -25.4048
16.0359 | -34.8721 | -25.4122
16.042 | -34.8726 | -25.4198
16.0481 | -34.8731 | -25.4276
16.0542 | -34.8737 | -25.4353
16.0603 | -34.8742 | -25.4428
16.0664 | -34.8747 | -25.4503
16.0725 | -34.8752 | -25.4578
16.0786 | -34.8757 | -25.4653
16.0847 | -34.8762 | -25.4731
16.0908 | -34.8767 | -25.4807
16.0969 | -34.8772 | -25.4886
16.103 | -34.8777 | -25.4963
16.1091 | -34.8782 | -25.5037
16.1152 | -34.8787 | -25.5114
16.1214 | -34.8792 | -25.5192
16.1275 | -34.8798 | -25.5268
16.1336 | -34.8803 | -25.5343
16.1397 | -34.8808 | -25.542
16.1458 | -34.8813 | -25.5495
16.1519 | -34.8818 | -25.5574
16.158 | -34.8823 | -25.5652
16.1641 | -34.8828 | -25.5729
16.1702 | -34.8833 | -25.5806
16.1763 | -34.8838 | -25.5883
16.1824 | -34.8843 | -25.5955
16.1885 | -34.8848 | -25.603
16.1946 | -34.8853 | -25.6105
16.2007 | -34.8858 | -25.6181
16.2068 | -34.8863 | -25.6256
16.2129 | -34.8868 | -25.6332
16.219 | -34.8873 | -25.6402
16.2251 | -34.8878 | -25.6478
16.2312 | -34.8883 | -25.6553
16.2373 | -34.8888 | -25.6627
16.2434 | -34.8893 | -25.6702
16.2495 | -34.8898 | -25.6777
16.2556 | -34.8904 | -25.6854
16.2618 | -34.8909 | -25.6929
16.2679 | -34.8914 | -25.7007
16.274 | -34.8919 | -25.7082
16.2801 | -34.8924 | -25.7157
16.2862 | -34.8929 | -25.723
16.2923 | -34.8934 | -25.7306
16.2984 | -34.894 | -25.7382
16.3045 | -34.8945 | -25.7459
16.3106 | -34.895 | -25.7536
16.3167 | -34.8955 | -25.761
16.3228 | -34.896 | -25.7683
16.3289 | -34.8965 | -25.7759
16.335 | -34.897 | -25.7832
16.3411 | -34.8975 | -25.7909
16.3472 | -34.898 | -25.7986
16.3533 | -34.8985 | -25.8061
16.3594 | -34.899 | -25.8135
16.3655 | -34.8995 | -25.8211
16.3716 | -34.9 | -25.8286
16.3777 | -34.9005 | -25.8362
16.3838 | -34.901 | -25.8437
16.3899 | -34.9015 | -25.8508
16.396 | -34.9021 | -25.8583
16.4021 | -34.9026 | -25.8659
16.4083 | -34.9031 | -25.8734
16.4144 | -34.9036 | -25.8808
16.4205 | -34.9041 | -25.8883
16.4266 | -34.9046 | -25.8955
16.4327 | -34.9051 | -25.903
16.4388 | -34.9056 | -25.9107
16.4449 | -34.9061 | -25.9182
16.451 | -34.9066 | -25.9258
16.4571 | -34.9071 | -25.9333
16.4632 | -34.9076 | -25.9403
16.4693 | -34.9081 | -25.9478
16.4754 | -34.9086 | -25.9554
16.4815 | -34.9091 | -25.9629
16.4876 | -34.9096 | -25.9704
16.4937 | -34.9101 | -25.9777
16.4998 | -34.9106 | -25.985
16.5059 | -34.9112 | -25.9923
16.512 | -34.9117 | -25.9999
16.5181 | -34.9122 | -26.0077
16.5242 | -34.9127 | -26.0153
16.5303 | -34.9132 | -26.0224
16.5364 | -34.9137 | -26.0298
16.5425 | -34.9142 | -26.0372
16.5487 | -34.9147 | -26.0446
16.5548 | -34.9152 | -26.0521
16.5609 | -34.9157 | -26.0596
16.567 | -34.9163 | -26.067
16.5731 | -34.9168 | -26.0746
16.5792 | -34.9173 | -26.0821
16.5853 | -34.9178 | -26.0897
16.5914 | -34.9183 | -26.0972
16.5975 | -34.9188 | -26.1048
16.6036 | -34.9193 | -26.112
16.6097 | -34.9198 | -26.1196
16.6158 | -34.9203 | -26.1269
16.6219 | -34.9208 | -26.1343
16.628 | -34.9213 | -26.1418
16.6341 | -34.9218 | -26.1492
16.6402 | -34.9223 | -26.1567
16.6463 | -34.9228 | -26.1642
16.6524 | -34.9233 | -26.1716
16.6585 | -34.9238 | -26.1792
16.6646 | -34.9244 | -26.1866
16.6707 | -34.9249 | -26.1938
16.6768 | -34.9254 | -26.2014
16.6829 | -34.9259 | -26.2087
16.689 | -34.9264 | -26.2163
16.6952 | -34.9269 | -26.2237
16.7013 | -34.9274 | -26.2311
16.7074 | -34.928 | -26.2384
16.7135 | -34.9285 | -26.246
16.7196 | -34.929 | -26.2533
16.7257 | -34.9296 | -26.2606
16.7318 | -34.9301 | -26.2678
16.7379 | -34.9306 | -26.2751
16.744 | -34.9311 | -26.2824
16.7501 | -34.9316 | -26.2896
16.7562 | -34.9321 | -26.2971
16.7623 | -34.9326 | -26.3043
16.7684 | -34.9331 | -26.3116
16.7745 | -34.9336 | -26.319
16.7806 | -34.9341 | -26.326
16.7867 | -34.9346 | -26.3334
16.7928 | -34.9351 | -26.3409
16.7989 | -34.9357 | -26.3483
16.805 | -34.9362 | -26.3558
16.8111 | -34.9367 | -26.363
16.8172 | -34.9372 | -26.3705
16.8233 | -34.9378 | -26.3778
16.8294 | -34.9383 | -26.3852
16.8356 | -34.9388 | -26.3927
16.8417 | -34.9393 | -26.4
16.8478 | -34.9398 | -26.4072
16.8539 | -34.9403 | -26.4145
16.86 | -34.9408 | -26.422
16.8661 | -34.9413 | -26.4293
16.8722 | -34.9418 | -26.4369
16.8783 | -34.9423 | -26.4444
16.8844 | -34.9428 | -26.4517
16.8905 | -34.9433 | -26.4592
16.8966 | -34.9439 | -26.4664
16.9027 | -34.9444 | -26.4738
16.9088 | -34.9449 | -26.4813
16.9149 | -34.9454 | -26.4887
16.921 | -34.9459 | -26.496
16.9271 | -34.9465 | -26.5035
16.9332 | -34.947 | -26.5108
16.9393 | -34.9475 | -26.5182
16.9454 | -34.948 | -26.5254
16.9515 | -34.9485 | -26.533
16.9576 | -34.949 | -26.54
16.9637 | -34.9495 | -26.5473
16.9698 | -34.95 | -26.5546
16.9759 | -34.9505 | -26.5619
16.9821 | -34.951 | -26.5695
16.9882 | -34.9515 | -26.5768
16.9943 | -34.9521 | -26.584
17.0004 | -34.9525 | -26.5913
17.0065 | -34.9531 | -26.5987
17.0126 | -34.9536 | -26.6061
17.0187 | -34.9541 | -26.6136
17.0248 | -34.9546 | -26.621
17.0309 | -34.9551 | -26.6284
17.037 | -34.9556 | -26.6358
17.0431 | -34.9561 | -26.6433
17.0492 | -34.9566 | -26.6506
17.0553 | -34.9571 | -26.658
17.0614 | -34.9576 | -26.6653
17.0675 | -34.9581 | -26.6727
17.0736 | -34.9587 | -26.6802
17.0797 | -34.9591 | -26.6875
17.0858 | -34.9597 | -26.6948
17.0919 | -34.9602 | -26.7022
17.098 | -34.9607 | -26.7094
17.1041 | -34.9612 | -26.7166
17.1102 | -34.9618 | -26.7241
17.1163 | -34.9623 | -26.7314
17.1225 | -34.9628 | -26.7387
17.1286 | -34.9633 | -26.7458
17.1347 | -34.9638 | -26.7531
17.1408 | -34.9644 | -26.7607
17.1469 | -34.9648 | -26.7679
17.153 | -34.9654 | -26.7754
17.1591 | -34.9658 | -26.7827
17.1652 | -34.9663 | -26.79
17.1713 | -34.9668 | -26.7972
17.1774 | -34.9674 | -26.8045
17.1835 | -34.9679 | -26.8116
17.1896 | -34.9683 | -26.8189
17.1957 | -34.9689 | -26.8264
17.2018 | -34.9694 | -26.8336
17.2079 | -34.9699 | -26.8409
17.214 | -34.9704 | -26.8483
17.2201 | -34.9709 | -26.8556
17.2262 | -34.9714 | -26.8631
17.2323 | -34.9719 | -26.8704
17.2384 | -34.9724 | -26.8776
17.2445 | -34.9729 | -26.8849
17.2506 | -34.9735 | -26.8923
17.2567 | -34.974 | -26.8995
17.2628 | -34.9745 | -26.9068
17.269 | -34.975 | -26.9139
17.2751 | -34.9755 | -26.9213
17.2812 | -34.976 | -26.9287
17.2873 | -34.9765 | -26.936
17.2934 | -34.977 | -26.9434
17.2995 | -34.9775 | -26.9508
17.3056 | -34.9781 | -26.9582
17.3117 | -34.9786 | -26.9655
17.3178 | -34.9791 | -26.9726
17.3239 | -34.9796 | -26.9799
17.33 | -34.9801 | -26.987
17.3361 | -34.9806 | -26.9943
17.3422 | -34.9811 | -27.0014
17.3483 | -34.9816 | -27.0087
17.3544 | -34.9821 | -27.0161
17.3605 | -34.9826 | -27.0233
17.3666 | -34.9831 | -27.0305
17.3727 | -34.9836 | -27.0378
17.3788 | -34.9841 | -27.0447
17.3849 | -34.9846 | -27.052
17.391 | -34.9851 | -27.0594
17.3971 | -34.9856 | -27.0667
17.4032 | -34.9861 | -27.074
17.4094 | -34.9866 | -27.0813
17.4155 | -34.9871 | -27.0887
17.4216 | -34.9876 | -27.096
17.4277 | -34.9882 | -27.1034
17.4338 | -34.9887 | -27.1108
17.4399 | -34.9892 | -27.1182
17.446 | -34.9897 | -27.1252
17.4521 | -34.9902 | -27.1325
17.4582 | -34.9907 | -27.1397
17.4643 | -34.9912 | -27.1469
17.4704 | -34.9917 | -27.1544
17.4765 | -34.9922 | -27.1618
17.4826 | -34.9927 | -27.169
17.4887 | -34.9932 | -27.1763
17.4948 | -34.9937 | -27.1836
17.5009 | -34.9942 | -27.1909
17.507 | -34.9947 | -27.1983
17.5131 | -34.9953 | -27.2057
17.5192 | -34.9957 | -27.2128
17.5253 | -34.9962 | -27.2202
17.5314 | -34.9967 | -27.2274
17.5375 | -34.9972 | -27.2347
17.5436 | -34.9977 | -27.242
17.5497 | -34.9982 | -27.2491
17.5559 | -34.9987 | -27.2561
17.562 | -34.9992 | -27.2633
17.5681 | -34.9998 | -27.2704
17.5742 | -35.0003 | -27.2778
17.5803 | -35.0008 | -27.2851
17.5864 | -35.0013 | -27.2924
17.5925 | -35.0018 | -27.2998
17.5986 | -35.0023 | -27.307
17.6047 | -35.0028 | -27.3143
17.6108 | -35.0033 | -27.3215
17.6169 | -35.0038 | -27.3287
17.623 | -35.0043 | -27.3358
17.6291 | -35.0048 | -27.3429
17.6352 | -35.0053 | -27.3502
17.6413 | -35.0058 | -27.3576
17.6474 | -35.0063 | -27.365
17.6535 | -35.0068 | -27.3723
17.6596 | -35.0073 | -27.3796
17.6657 | -35.0078 | -27.3868
17.6718 | -35.0083 | -27.394
17.6779 | -35.0088 | -27.4012
17.684 | -35.0093 | -27.4084
17.6901 | -35.0098 | -27.4158
17.6963 | -35.0103 | -27.4229
17.7024 | -35.0108 | -27.4301
17.7085 | -35.0113 | -27.4373
17.7146 | -35.0119 | -27.4444
17.7207 | -35.0124 | -27.4515
17.7268 | -35.0129 | -27.4587
17.7329 | -35.0134 | -27.466
17.739 | -35.0139 | -27.4732
17.7451 | -35.0144 | -27.4803
17.7512 | -35.0149 | -27.4874
17.7573 | -35.0154 | -27.4947
17.7634 | -35.0159 | -27.502
17.7695 | -35.0164 | -27.5091
17.7756 | -35.0169 | -27.5164
17.7817 | -35.0174 | -27.5236
17.7878 | -35.0179 | -27.5309
17.7939 | -35.0184 | -27.5381
17.8 | -35.0189 | -27.5455
17.8061 | -35.0194 | -27.5526
17.8122 | -35.0199 | -27.5599
17.8183 | -35.0204 | -27.5672
17.8244 | -35.0209 | -27.5744
17.8305 | -35.0214 | -27.5818
17.8366 | -35.0219 | -27.5889
17.8428 | -35.0224 | -27.5962
17.8489 | -35.0229 | -27.6033
17.855 | -35.0234 | -27.6105
17.8611 | -35.0239 | -27.6177
17.8672 | -35.0244 | -27.6249
17.8733 | -35.0249 | -27.6322
17.8794 | -35.0254 | -27.6394
17.8855 | -35.0259 | -27.6467
17.8916 | -35.0264 | -27.6539
17.8977 | -35.0269 | -27.6611
17.9038 | -35.0274 | -27.668
17.9099 | -35.0279 | -27.6753
17.916 | -35.0284 | -27.6825
17.9221 | -35.0289 | -27.6897
17.9282 | -35.0294 | -27.6967
17.9343 | -35.0299 | -27.704
17.9404 | -35.0304 | -27.7111
17.9465 | -35.0309 | -27.7183
17.9526 | -35.0314 | -27.7256
17.9587 | -35.0319 | -27.7329
17.9648 | -35.0324 | -27.7401
17.9709 | -35.0329 | -27.7473
17.977 | -35.0334 | -27.7546
17.9832 | -35.0338 | -27.7619
17.9893 | -35.0343 | -27.7693
17.9954 | -35.0348 | -27.7766
18.0015 | -35.0353 | -27.7839
18.0076 | -35.0358 | -27.791
18.0137 | -35.0363 | -27.7982
18.0198 | -35.0368 | -27.8053
18.0259 | -35.0373 | -27.8125
18.032 | -35.0378 | -27.8198
18.0381 | -35.0382 | -27.8269
18.0442 | -35.0387 | -27.8342
18.0503 | -35.0392 | -27.8416
18.0564 | -35.0397 | -27.8486
18.0625 | -35.0402 | -27.8559
18.0686 | -35.0407 | -27.863
18.0747 | -35.0412 | -27.8703
18.0808 | -35.0417 | -27.8774
18.0869 | -35.0422 | -27.8846
18.093 | -35.0427 | -27.8918
18.0991 | -35.0432 | -27.899
18.1052 | -35.0437 | -27.9063
18.1113 | -35.0442 | -27.9135
18.1174 | -35.0447 | -27.9206
18.1236 | -35.0452 | -27.9279
18.1297 | -35.0457 | -27.9352
18.1358 | -35.0462 | -27.9424
18.1419 | -35.0467 | -27.9498
18.148 | -35.0472 | -27.9569
18.1541 | -35.0477 | -27.9639
18.1602 | -35.0482 | -27.9712
18.1663 | -35.0487 | -27.9783
18.1724 | -35.0492 | -27.9854
18.1785 | -35.0497 | -27.9926
18.1846 | -35.0501 | -27.9999
18.1907 | -35.0506 | -28.0069
18.1968 | -35.0511 | -28.0143
18.2029 | -35.0516 | -28.0216
18.209 | -35.0521 | -28.0289
18.2151 | -35.0525 | -28.0362
18.2212 | -35.0531 | -28.0432
18.2273 | -35.0535 | -28.0504
18.2334 | -35.0541 | -28.0576
18.2395 | -35.0545 | -28.0647
18.2456 | -35.055 | -28.0719
18.2517 | -35.0555 | -28.0791
18.2578 | -35.056 | -28.0864
18.2639 | -35.0565 | -28.0938
18.2701 | -35.057 | -28.1011
18.2762 | -35.0575 | -28.1082
18.2823 | -35.058 | -28.1154
18.2884 | -35.0585 | -28.1224
18.2945 | -35.059 | -28.1293
18.3006 | -35.0595 | -28.1365
18.3067 | -35.06 | -28.1435
18.3128 | -35.0605 | -28.1507
18.3189 | -35.061 | -28.1578
18.325 | -35.0615 | -28.1648
18.3311 | -35.0619 | -28.1722
18.3372 | -35.0625 | -28.1793
18.3433 | -35.063 | -28.1866
18.3494 | -35.0634 | -28.1938
18.3555 | -35.064 | -28.2009
18.3616 | -35.0645 | -28.2078
18.3677 | -35.065 | -28.2148
18.3738 | -35.0655 | -28.2218
18.3799 | -35.066 | -28.2289
18.386 | -35.0664 | -28.2359
18.3921 | -35.0669 | -28.243
18.3982 | -35.0674 | -28.25
18.4043 | -35.0679 | -28.2571
18.4105 | -35.0683 | -28.2641
18.4166 | -35.0688 | -28.2711
18.4227 | -35.0693 | -28.2784
18.4288 | -35.0698 | -28.2856
18.4349 | -35.0703 | -28.2928
18.441 | -35.0708 | -28.2999
18.4471 | -35.0713 | -28.3069
18.4532 | -35.0717 | -28.314
18.4593 | -35.0722 | -28.3213
18.4654 | -35.0727 | -28.3284
18.4715 | -35.0732 | -28.3355
18.4776 | -35.0736 | -28.3429
18.4837 | -35.0741 | -28.3502
18.4898 | -35.0746 | -28.3574
18.4959 | -35.0751 | -28.3646
18.502 | -35.0755 | -28.3718
18.5081 | -35.076 | -28.3789
18.5142 | -35.0765 | -28.3861
18.5203 | -35.077 | -28.3932
18.5264 | -35.0775 | -28.4004
18.5325 | -35.078 | -28.4076
18.5386 | -35.0785 | -28.4146
18.5447 | -35.0789 | -28.4217
18.5508 | -35.0794 | -28.4288
18.557 | -35.0799 | -28.436
18.5631 | -35.0804 | -28.443
18.5692 | -35.0809 | -28.4501
18.5753 | -35.0813 | -28.4573
18.5814 | -35.0818 | -28.4643
18.5875 | -35.0823 | -28.4715
18.5936 | -35.0828 | -28.4785
18.5997 | -35.0833 | -28.4856
18.6058 | -35.0837 | -28.4927
18.6119 | -35.0842 | -28.4998
18.618 | -35.0847 | -28.5072
18.6241 | -35.0852 | -28.5143
18.6302 | -35.0857 | -28.5216
18.6363 | -35.0862 | -28.5288
18.6424 | -35.0866 | -28.536
18.6485 | -35.0871 | -28.5432
18.6546 | -35.0876 | -28.5503
18.6607 | -35.0881 | -28.5574
18.6668 | -35.0886 | -28.5647
18.6729 | -35.0891 | -28.572
18.679 | -35.0895 | -28.5789
18.6851 | -35.09 | -28.5859
18.6912 | -35.0905 | -28.5929
18.6974 | -35.091 | -28.6
18.7035 | -35.0914 | -28.6072
18.7096 | -35.092 | -28.6145
18.7157 | -35.0925 | -28.6218
18.7218 | -35.093 | -28.629
18.7279 | -35.0935 | -28.6361
18.734 | -35.094 | -28.6432
18.7401 | -35.0945 | -28.6504
18.7462 | -35.095 | -28.6571
18.7523 | -35.0955 | -28.6644
18.7584 | -35.096 | -28.6715
18.7645 | -35.0965 | -28.6786
18.7706 | -35.0969 | -28.6856
18.7767 | -35.0974 | -28.6927
18.7828 | -35.0979 | -28.6995
18.7889 | -35.0984 | -28.7068
18.795 | -35.0989 | -28.7143
18.8011 | -35.0993 | -28.7214
18.8072 | -35.0998 | -28.7286
18.8133 | -35.1003 | -28.7358
18.8194 | -35.1007 | -28.7425
18.8255 | -35.1012 | -28.7495
18.8316 | -35.1017 | -28.7566
18.8377 | -35.1022 | -28.7634
18.8439 | -35.1027 | -28.7703
18.85 | -35.1031 | -28.7773
18.8561 | -35.1036 | -28.7846
18.8622 | -35.1041 | -28.7916
18.8683 | -35.1046 | -28.7986
18.8744 | -35.105 | -28.8057
18.8805 | -35.1055 | -28.8128
18.8866 | -35.106 | -28.8201
18.8927 | -35.1065 | -28.8269
18.8988 | -35.107 | -28.8339
18.9049 | -35.1075 | -28.841
18.911 | -35.108 | -28.8482
18.9171 | -35.1085 | -28.8554
18.9232 | -35.109 | -28.8624
18.9293 | -35.1095 | -28.8695
18.9354 | -35.1099 | -28.8766
18.9415 | -35.1104 | -28.8836
18.9476 | -35.1109 | -28.8905
18.9537 | -35.1114 | -28.8977
18.9598 | -35.1119 | -28.9047
18.9659 | -35.1124 | -28.9117
18.972 | -35.1129 | -28.9187
18.9781 | -35.1134 | -28.9259
18.9843 | -35.1138 | -28.9332
18.9904 | -35.1143 | -28.9403
18.9965 | -35.1148 | -28.9469
19.0026 | -35.1152 | -28.9542
19.0087 | -35.1157 | -28.9612
19.0148 | -35.1162 | -28.9682
19.0209 | -35.1167 | -28.9755
19.027 | -35.1172 | -28.9827
19.0331 | -35.1177 | -28.9897
19.0392 | -35.1182 | -28.9967
19.0453 | -35.1186 | -29.0039
19.0514 | -35.1191 | -29.0109
19.0575 | -35.1196 | -29.018
19.0636 | -35.12 | -29.0249
19.0697 | -35.1205 | -29.0319
19.0758 | -35.121 | -29.0388
19.0819 | -35.1215 | -29.0459
19.088 | -35.122 | -29.0529
19.0941 | -35.1225 | -29.0601
19.1002 | -35.1229 | -29.067
19.1063 | -35.1234 | -29.0742
19.1124 | -35.1239 | -29.0811
19.1185 | -35.1244 | -29.0883
19.1246 | -35.1249 | -29.0955
19.1308 | -35.1254 | -29.1025
19.1369 | -35.1259 | -29.1093
19.143 | -35.1264 | -29.1164
19.1491 | -35.1268 | -29.1236
19.1552 | -35.1273 | -29.1308
19.1613 | -35.1277 | -29.1377
19.1674 | -35.1282 | -29.1447
19.1735 | -35.1287 | -29.1516
19.1796 | -35.1292 | -29.1586
19.1857 | -35.1296 | -29.1657
19.1918 | -35.1301 | -29.1726
19.1979 | -35.1306 | -29.1799
19.204 | -35.1311 | -29.1868
19.2101 | -35.1315 | -29.1936
19.2162 | -35.132 | -29.2007
19.2223 | -35.1325 | -29.2076
19.2284 | -35.1329 | -29.2143
19.2345 | -35.1334 | -29.2215
19.2406 | -35.1339 | -29.2287
19.2467 | -35.1344 | -29.2359
19.2528 | -35.1349 | -29.243
19.2589 | -35.1353 | -29.2501
19.265 | -35.1358 | -29.2572
19.2712 | -35.1362 | -29.2641
19.2773 | -35.1367 | -29.2711
19.2834 | -35.1372 | -29.278
19.2895 | -35.1377 | -29.285
19.2956 | -35.1382 | -29.2919
19.3017 | -35.1386 | -29.2987
19.3078 | -35.1391 | -29.3057
19.3139 | -35.1396 | -29.3125
19.32 | -35.1401 | -29.3194
19.3261 | -35.1406 | -29.3265
19.3322 | -35.1411 | -29.3332
19.3383 | -35.1415 | -29.3399
19.3444 | -35.142 | -29.3465
19.3505 | -35.1425 | -29.3534
19.3566 | -35.143 | -29.3603
19.3627 | -35.1434 | -29.3672
19.3688 | -35.1439 | -29.3739
19.3749 | -35.1444 | -29.381
19.381 | -35.1449 | -29.3878
19.3871 | -35.1453 | -29.3947
19.3932 | -35.1458 | -29.4018
19.3993 | -35.1463 | -29.4087
19.4054 | -35.1467 | -29.4156
19.4115 | -35.1472 | -29.4226
19.4177 | -35.1477 | -29.4295
19.4238 | -35.1482 | -29.4367
19.4299 | -35.1486 | -29.4435
19.436 | -35.1492 | -29.4505
19.4421 | -35.1496 | -29.4575
19.4482 | -35.1501 | -29.4645
19.4543 | -35.1506 | -29.4709
19.4604 | -35.151 | -29.4779
19.4665 | -35.1515 | -29.4849
19.4726 | -35.1519 | -29.4919
19.4787 | -35.1524 | -29.4988
19.4848 | -35.1528 | -29.5056
19.4909 | -35.1534 | -29.5126
19.497 | -35.1538 | -29.5195
19.5031 | -35.1543 | -29.5268
19.5092 | -35.1547 | -29.5336
19.5153 | -35.1552 | -29.5401
19.5214 | -35.1557 | -29.5475
19.5275 | -35.1562 | -29.5544
19.5336 | -35.1566 | -29.5614
19.5397 | -35.1572 | -29.5682
19.5458 | -35.1577 | -29.5754
19.5519 | -35.1581 | -29.5824
19.5581 | -35.1585 | -29.5893
19.5642 | -35.159 | -29.5962
19.5703 | -35.1594 | -29.603
19.5764 | -35.1599 | -29.6098
19.5825 | -35.1604 | -29.6166
19.5886 | -35.1609 | -29.6236
19.5947 | -35.1614 | -29.6303
19.6008 | -35.1619 | -29.6373
19.6069 | -35.1624 | -29.6438
19.613 | -35.1628 | -29.6509
19.6191 | -35.1633 | -29.6574
19.6252 | -35.1637 | -29.6643
19.6313 | -35.1642 | -29.6711
19.6374 | -35.1646 | -29.6779
19.6435 | -35.1651 | -29.6847
19.6496 | -35.1656 | -29.6916
19.6557 | -35.1661 | -29.6983
19.6618 | -35.1665 | -29.705
19.6679 | -35.1669 | -29.7119
19.674 | -35.1674 | -29.7185
19.6801 | -35.1679 | -29.7254
19.6862 | -35.1684 | -29.7322
19.6923 | -35.1689 | -29.7392
19.6984 | -35.1694 | -29.7453
19.7046 | -35.1699 | -29.7524
19.7107 | -35.1703 | -29.7593
19.7168 | -35.1708 | -29.7662
19.7229 | -35.1713 | -29.7731
19.729 | -35.1717 | -29.78
19.7351 | -35.1722 | -29.7868
19.7412 | -35.1727 | -29.7939
19.7473 | -35.1732 | -29.8006
19.7534 | -35.1738 | -29.8074
19.7595 | -35.1742 | -29.8141
19.7656 | -35.1746 | -29.821
19.7717 | -35.1751 | -29.8275
19.7778 | -35.1756 | -29.8347
19.7839 | -35.176 | -29.8415
19.79 | -35.1765 | -29.8484
19.7961 | -35.1769 | -29.8548
19.8022 | -35.1773 | -29.8616
19.8083 | -35.1778 | -29.8679
19.8144 | -35.1783 | -29.8751
19.8205 | -35.1788 | -29.8818
19.8266 | -35.1793 | -29.8888
19.8327 | -35.1797 | -29.8952
19.8388 | -35.1801 | -29.902
19.845 | -35.1805 | -29.9089
19.8511 | -35.1809 | -29.9159
19.8572 | -35.1814 | -29.9225
19.8633 | -35.1819 | -29.9291
19.8694 | -35.1823 | -29.936
19.8755 | -35.1829 | -29.9422
19.8816 | -35.1833 | -29.949
19.8877 | -35.1838 | -29.9561
19.8938 | -35.1842 | -29.9629
19.8999 | -35.1847 | -29.9699
19.906 | -35.1851 | -29.9768
19.9121 | -35.1856 | -29.9832
19.9182 | -35.1861 | -29.9901
19.9243 | -35.1865 | -29.9969
19.9304 | -35.1869 | -30.0038
19.9365 | -35.1874 | -30.0107
19.9426 | -35.1879 | -30.0174
19.9487 | -35.1884 | -30.0245
19.9548 | -35.1888 | -30.0313
19.9609 | -35.1893 | -30.0379
19.967 | -35.1897 | -30.0448
19.9731 | -35.1902 | -30.0517
19.9792 | -35.1907 | -30.0582
19.9853 | -35.1912 | -30.0647
19.9915 | -35.1917 | -30.0713
19.9976 | -35.1922 | -30.0781
20.0037 | -35.1927 | -30.085
20.0098 | -35.1931 | -30.0918
20.0159 | -35.1936 | -30.0987
20.022 | -35.194 | -30.1054
20.0281 | -35.1945 | -30.1122
20.0342 | -35.1949 | -30.119
20.0403 | -35.1954 | -30.1258
20.0464 | -35.1958 | -30.1324
20.0525 | -35.1963 | -30.1391
20.0586 | -35.1968 | -30.1459
20.0647 | -35.1972 | -30.1528
20.0708 | -35.1977 | -30.16
20.0769 | -35.1981 | -30.167
20.083 | -35.1986 | -30.1738
20.0891 | -35.1991 | -30.1806
20.0952 | -35.1995 | -30.1872
20.1013 | -35.1999 | -30.1942
20.1074 | -35.2004 | -30.2007
20.1135 | -35.2009 | -30.2075
20.1196 | -35.2013 | -30.2138
20.1257 | -35.2017 | -30.2206
20.1319 | -35.2022 | -30.2273
20.138 | -35.2027 | -30.2339
20.1441 | -35.2031 | -30.2408
20.1502 | -35.2036 | -30.2473
20.1563 | -35.204 | -30.2536
20.1624 | -35.2044 | -30.2602
20.1685 | -35.2049 | -30.2673
20.1746 | -35.2054 | -30.2741
20.1807 | -35.2058 | -30.2806
20.1868 | -35.2064 | -30.2875
20.1929 | -35.2068 | -30.2939
20.199 | -35.2073 | -30.3008
20.2051 | -35.2077 | -30.3074
20.2112 | -35.2082 | -30.314
20.2173 | -35.2087 | -30.3208
20.2234 | -35.2091 | -30.3275
20.2295 | -35.2096 | -30.3343
20.2356 | -35.2101 | -30.3413
20.2417 | -35.2105 | -30.3478
20.2478 | -35.2109 | -30.3543
20.2539 | -35.2114 | -30.3613
20.26 | -35.2118 | -30.3676
20.2661 | -35.2123 | -30.3746
20.2723 | -35.2127 | -30.3812
20.2784 | -35.2132 | -30.388
20.2845 | -35.2136 | -30.3948
20.2906 | -35.2141 | -30.4017
20.2967 | -35.2145 | -30.4079
20.3028 | -35.215 | -30.4148
20.3089 | -35.2154 | -30.4213
20.315 | -35.2159 | -30.4279
20.3211 | -35.2164 | -30.4346
20.3272 | -35.2168 | -30.4412
20.3333 | -35.2173 | -30.4479
20.3394 | -35.2178 | -30.4544
20.3455 | -35.2183 | -30.461
20.3516 | -35.2187 | -30.4675
20.3577 | -35.2192 | -30.4743
20.3638 | -35.2196 | -30.4811
20.3699 | -35.2201 | -30.4876
20.376 | -35.2207 | -30.4942
20.3821 | -35.2211 | -30.5009
20.3882 | -35.2216 | -30.5077
20.3943 | -35.222 | -30.5144
20.4004 | -35.2225 | -30.5208
20.4065 | -35.2229 | -30.5271
20.4126 | -35.2234 | -30.5339
20.4188 | -35.2238 | -30.5405
20.4249 | -35.2243 | -30.5469
20.431 | -35.2247 | -30.5534
20.4371 | -35.2252 | -30.5599
20.4432 | -35.2256 | -30.5665
20.4493 | -35.2261 | -30.5733
20.4554 | -35.2266 | -30.58
20.4615 | -35.227 | -30.5866
20.4676 | -35.2275 | -30.5936
20.4737 | -35.228 | -30.5999
20.4798 | -35.2284 | -30.6065
20.4859 | -35.2289 | -30.6134
20.492 | -35.2293 | -30.62
20.4981 | -35.2298 | -30.6267
20.5042 | -35.2303 | -30.6332
20.5103 | -35.2307 | -30.6394
20.5164 | -35.2312 | -30.6461
20.5225 | -35.2316 | -30.6527
20.5286 | -35.2321 | -30.6592
20.5347 | -35.2325 | -30.6656
20.5408 | -35.233 | -30.6723
20.5469 | -35.2335 | -30.6791
20.553 | -35.2339 | -30.6859
20.5592 | -35.2344 | -30.6926
20.5653 | -35.2348 | -30.6991
20.5714 | -35.2353 | -30.7058
20.5775 | -35.2358 | -30.7122
20.5836 | -35.2362 | -30.7189
20.5897 | -35.2366 | -30.7257
20.5958 | -35.2371 | -30.7325
20.6019 | -35.2376 | -30.7392
20.608 | -35.238 | -30.7461
20.6141 | -35.2384 | -30.7526
20.6202 | -35.2389 | -30.7591
20.6263 | -35.2394 | -30.7657
20.6324 | -35.2399 | -30.7724
20.6385 | -35.2403 | -30.7792
20.6446 | -35.2408 | -30.7859
20.6507 | -35.2412 | -30.7923
20.6568 | -35.2417 | -30.7988
20.6629 | -35.2422 | -30.8054
20.669 | -35.2426 | -30.8122
20.6751 | -35.2431 | -30.8188
20.6812 | -35.2435 | -30.8251
20.6873 | -35.2439 | -30.8319
20.6934 | -35.2443 | -30.8387
20.6995 | -35.2447 | -30.8453
20.7057 | -35.2452 | -30.852
20.7118 | -35.2457 | -30.8585
20.7179 | -35.2461 | -30.8648
20.724 | -35.2466 | -30.8712
20.7301 | -35.247 | -30.8776
20.7362 | -35.2475 | -30.8843
20.7423 | -35.248 | -30.8908
20.7484 | -35.2485 | -30.8971
20.7545 | -35.249 | -30.9035
20.7606 | -35.2495 | -30.9102
20.7667 | -35.2499 | -30.9168
20.7728 | -35.2504 | -30.923
20.7789 | -35.2509 | -30.9297
20.785 | -35.2513 | -30.9365
20.7911 | -35.2517 | -30.943
20.7972 | -35.2522 | -30.9496
20.8033 | -35.2526 | -30.9565
20.8094 | -35.2531 | -30.963
20.8155 | -35.2535 | -30.9698
20.8216 | -35.254 | -30.9766
20.8277 | -35.2544 | -30.9827
20.8338 | -35.2548 | -30.9893
20.8399 | -35.2553 | -30.996
20.8461 | -35.2558 | -31.0026
20.8522 | -35.2562 | -31.0096
20.8583 | -35.2567 | -31.0157
20.8644 | -35.2571 | -31.0223
20.8705 | -35.2575 | -31.0291
20.8766 | -35.258 | -31.036
20.8827 | -35.2584 | -31.0424
20.8888 | -35.2588 | -31.0491
20.8949 | -35.2593 | -31.0556
20.901 | -35.2597 | -31.0621
20.9071 | -35.2602 | -31.0683
20.9132 | -35.2607 | -31.0745
20.9193 | -35.2612 | -31.0812
20.9254 | -35.2616 | -31.0878
20.9315 | -35.2621 | -31.094
20.9376 | -35.2626 | -31.1007
20.9437 | -35.263 | -31.1071
20.9498 | -35.2634 | -31.1134
20.9559 | -35.2639 | -31.12
20.962 | -35.2643 | -31.1265
20.9681 | -35.2647 | -31.1325
20.9742 | -35.2652 | -31.139
20.9803 | -35.2656 | -31.1456
20.9864 | -35.266 | -31.1524
20.9926 | -35.2665 | -31.1589
20.9987 | -35.2669 | -31.1653
21.0048 | -35.2675 | -31.1718
21.0109 | -35.2679 | -31.1783
21.017 | -35.2684 | -31.1847
21.0231 | -35.2689 | -31.1911
21.0292 | -35.2693 | -31.1978
21.0353 | -35.2697 | -31.2041
21.0414 | -35.2702 | -31.2106
21.0475 | -35.2706 | -31.2173
21.0536 | -35.2711 | -31.2238
21.0597 | -35.2715 | -31.2302
21.0658 | -35.272 | -31.2366
21.0719 | -35.2724 | -31.2434
21.078 | -35.2728 | -31.2498
21.0841 | -35.2733 | -31.2563
21.0902 | -35.2737 | -31.263
21.0963 | -35.2741 | -31.2697
21.1024 | -35.2746 | -31.2757
21.1085 | -35.275 | -31.282
21.1146 | -35.2755 | -31.2885
21.1207 | -35.276 | -31.2951
21.1268 | -35.2764 | -31.3014
21.133 | -35.2769 | -31.3078
21.1391 | -35.2774 | -31.314
21.1452 | -35.2778 | -31.3203
21.1513 | -35.2782 | -31.3266
21.1574 | -35.2787 | -31.3333
21.1635 | -35.2792 | -31.3398
21.1696 | -35.2796 | -31.3463
21.1757 | -35.2801 | -31.3527
21.1818 | -35.2805 | -31.3592
21.1879 | -35.2809 | -31.3656
21.194 | -35.2814 | -31.3721
21.2001 | -35.2818 | -31.3782
21.2062 | -35.2823 | -31.3844
21.2123 | -35.2827 | -31.3909
21.2184 | -35.2831 | -31.3973
21.2245 | -35.2835 | -31.4037
21.2306 | -35.284 | -31.4101
21.2367 | -35.2845 | -31.4165
21.2428 | -35.2849 | -31.423
21.2489 | -35.2853 | -31.4294
21.255 | -35.2858 | -31.4362
21.2611 | -35.2862 | -31.4427
21.2672 | -35.2866 | -31.4492
21.2733 | -35.2871 | -31.4557
21.2795 | -35.2876 | -31.4617
21.2856 | -35.288 | -31.4682
21.2917 | -35.2884 | -31.4749
21.2978 | -35.2889 | -31.4811
21.3039 | -35.2893 | -31.4875
21.31 | -35.2898 | -31.4943
21.3161 | -35.2902 | -31.5005
21.3222 | -35.2907 | -31.5072
21.3283 | -35.2911 | -31.5139
21.3344 | -35.2916 | -31.5203
21.3405 | -35.292 | -31.5269
21.3466 | -35.2925 | -31.5332
21.3527 | -35.2929 | -31.5397
21.3588 | -35.2934 | -31.5461
21.3649 | -35.2938 | -31.5524
21.371 | -35.2943 | -31.5592
21.3771 | -35.2948 | -31.5658
21.3832 | -35.2952 | -31.5725
21.3893 | -35.2957 | -31.5792
21.3954 | -35.2961 | -31.5856
21.4015 | -35.2966 | -31.5919
21.4076 | -35.2971 | -31.5985
21.4137 | -35.2975 | -31.6048
21.4199 | -35.298 | -31.6114
21.426 | -35.2984 | -31.6179
21.4321 | -35.2988 | -31.6242
21.4382 | -35.2992 | -31.631
21.4443 | -35.2996 | -31.6373
21.4504 | -35.3001 | -31.6437
21.4565 | -35.3005 | -31.65
21.4626 | -35.3009 | -31.6565
21.4687 | -35.3013 | -31.6629
21.4748 | -35.3018 | -31.6693
21.4809 | -35.3022 | -31.6755
21.487 | -35.3027 | -31.6819
21.4931 | -35.3031 | -31.6883
21.4992 | -35.3036 | -31.6949
21.5053 | -35.304 | -31.7013
21.5114 | -35.3045 | -31.7077
21.5175 | -35.3049 | -31.7141
21.5236 | -35.3054 | -31.7207
21.5297 | -35.3058 | -31.7271
21.5358 | -35.3062 | -31.7337
21.5419 | -35.3066 | -31.7403
21.548 | -35.307 | -31.7468
21.5541 | -35.3075 | -31.753
21.5602 | -35.3079 | -31.7597
21.5664 | -35.3084 | -31.7659
21.5725 | -35.3088 | -31.7722
21.5786 | -35.3093 | -31.7789
21.5847 | -35.3097 | -31.7855
21.5908 | -35.3101 | -31.7918
21.5969 | -35.3106 | -31.7979
21.603 | -35.311 | -31.8042
21.6091 | -35.3114 | -31.811
21.6152 | -35.3118 | -31.8176
21.6213 | -35.3123 | -31.824
21.6274 | -35.3127 | -31.8304
21.6335 | -35.3131 | -31.8368
21.6396 | -35.3135 | -31.8429
21.6457 | -35.314 | -31.8492
21.6518 | -35.3144 | -31.8553
21.6579 | -35.3148 | -31.8618
21.664 | -35.3152 | -31.8684
21.6701 | -35.3156 | -31.8747
21.6762 | -35.316 | -31.8809
21.6823 | -35.3165 | -31.8874
21.6884 | -35.3169 | -31.8939
21.6945 | -35.3173 | -31.9004
21.7006 | -35.3177 | -31.907
21.7068 | -35.3182 | -31.9132
21.7129 | -35.3187 | -31.9194
21.719 | -35.3191 | -31.9259
21.7251 | -35.3195 | -31.9321
21.7312 | -35.3199 | -31.9384
21.7373 | -35.3204 | -31.9445
21.7434 | -35.3208 | -31.9511
21.7495 | -35.3212 | -31.9577
21.7556 | -35.3216 | -31.9641
21.7617 | -35.3221 | -31.9703
21.7678 | -35.3225 | -31.9767
21.7739 | -35.3229 | -31.9825
21.78 | -35.3233 | -31.9887
21.7861 | -35.3237 | -31.9953
21.7922 | -35.3242 | -32.0019
21.7983 | -35.3246 | -32.0082
21.8044 | -35.325 | -32.0144
21.8105 | -35.3254 | -32.0203
21.8166 | -35.3259 | -32.0268
21.8227 | -35.3263 | -32.0332
21.8288 | -35.3267 | -32.0393
21.8349 | -35.3272 | -32.0457
21.841 | -35.3276 | -32.0521
21.8471 | -35.3281 | -32.0581
21.8533 | -35.3286 | -32.0648
21.8594 | -35.329 | -32.0712
21.8655 | -35.3294 | -32.0776
21.8716 | -35.3298 | -32.0838
21.8777 | -35.3303 | -32.0904
21.8838 | -35.3307 | -32.0966
21.8899 | -35.3311 | -32.1029
21.896 | -35.3316 | -32.1092
21.9021 | -35.332 | -32.1156
21.9082 | -35.3325 | -32.1221
21.9143 | -35.3329 | -32.128
21.9204 | -35.3333 | -32.1342
21.9265 | -35.3337 | -32.1403
21.9326 | -35.3341 | -32.1464
21.9387 | -35.3346 | -32.153
21.9448 | -35.335 | -32.1593
21.9509 | -35.3355 | -32.1652
21.957 | -35.3359 | -32.1715
21.9631 | -35.3363 | -32.1778
21.9692 | -35.3367 | -32.1843
21.9753 | -35.3371 | -32.1906
21.9814 | -35.3376 | -32.197
21.9875 | -35.338 | -32.203
21.9937 | -35.3384 | -32.2092
21.9998 | -35.3388 | -32.2155
22.0059 | -35.3392 | -32.2218
22.012 | -35.3396 | -32.2279
22.0181 | -35.3401 | -32.2341
22.0242 | -35.3405 | -32.2404
22.0303 | -35.3409 | -32.2469
22.0364 | -35.3414 | -32.2533
22.0425 | -35.3418 | -32.2596
22.0486 | -35.3423 | -32.266
22.0547 | -35.3427 | -32.2722
22.0608 | -35.3431 | -32.2783
22.0669 | -35.3435 | -32.2845
22.073 | -35.344 | -32.2908
22.0791 | -35.3444 | -32.2971
22.0852 | -35.3449 | -32.3036
22.0913 | -35.3453 | -32.3097
22.0974 | -35.3457 | -32.3161
22.1035 | -35.3461 | -32.3226
22.1096 | -35.3466 | -32.3291
22.1157 | -35.347 | -32.3357
22.1218 | -35.3475 | -32.3421
22.1279 | -35.3479 | -32.3483
22.134 | -35.3483 | -32.3546
22.1402 | -35.3488 | -32.3608
22.1463 | -35.3492 | -32.3671
22.1524 | -35.3496 | -32.3734
22.1585 | -35.35 | -32.3795
22.1646 | -35.3504 | -32.3857
22.1707 | -35.3508 | -32.392
22.1768 | -35.3513 | -32.3982
22.1829 | -35.3517 | -32.4043
22.189 | -35.3522 | -32.4107
22.1951 | -35.3525 | -32.4167
22.2012 | -35.353 | -32.4228
22.2073 | -35.3534 | -32.4291
22.2134 | -35.3538 | -32.435
22.2195 | -35.3543 | -32.4411
22.2256 | -35.3547 | -32.4472
22.2317 | -35.3552 | -32.4534
22.2378 | -35.3556 | -32.4598
22.2439 | -35.356 | -32.4661
22.25 | -35.3564 | -32.4723
22.2561 | -35.3568 | -32.4784
22.2622 | -35.3572 | -32.4846
22.2683 | -35.3576 | -32.4909
22.2744 | -35.3581 | -32.4975
22.2806 | -35.3585 | -32.5037
22.2867 | -35.359 | -32.5099
22.2928 | -35.3594 | -32.5161
22.2989 | -35.3598 | -32.5222
22.305 | -35.3603 | -32.5283
22.3111 | -35.3607 | -32.5345
22.3172 | -35.3612 | -32.5407
22.3233 | -35.3616 | -32.5467
22.3294 | -35.3621 | -32.5531
22.3355 | -35.3625 | -32.5589
22.3416 | -35.3629 | -32.5652
22.3477 | -35.3633 | -32.5715
22.3538 | -35.3638 | -32.5779
22.3599 | -35.3642 | -32.5842
22.366 | -35.3646 | -32.5902
22.3721 | -35.365 | -32.5961
22.3782 | -35.3654 | -32.6025
22.3843 | -35.3659 | -32.6088
22.3904 | -35.3663 | -32.6149
22.3965 | -35.3667 | -32.6213
22.4026 | -35.3671 | -32.6276
22.4087 | -35.3676 | -32.6334
22.4148 | -35.368 | -32.6396
22.4209 | -35.3684 | -32.6457
22.4271 | -35.3689 | -32.6517
22.4332 | -35.3694 | -32.658
22.4393 | -35.3698 | -32.6643
22.4454 | -35.3703 | -32.67
22.4515 | -35.3707 | -32.6761
22.4576 | -35.3711 | -32.6823
22.4637 | -35.3715 | -32.6886
22.4698 | -35.372 | -32.6949
22.4759 | -35.3723 | -32.7008
22.482 | -35.3728 | -32.7072
22.4881 | -35.3732 | -32.7133
22.4942 | -35.3736 | -32.7197
22.5003 | -35.374 | -32.726
22.5064 | -35.3745 | -32.7321
22.5125 | -35.3749 | -32.7381
22.5186 | -35.3753 | -32.7443
22.5247 | -35.3757 | -32.7505
22.5308 | -35.3762 | -32.757
22.5369 | -35.3766 | -32.7631
22.543 | -35.377 | -32.7693
22.5491 | -35.3774 | -32.7753
22.5552 | -35.3778 | -32.7814
22.5613 | -35.3783 | -32.7878
22.5675 | -35.3787 | -32.7939
22.5736 | -35.3791 | -32.8002
22.5797 | -35.3795 | -32.8066
22.5858 | -35.3799 | -32.8126
22.5919 | -35.3804 | -32.8187
22.598 | -35.3808 | -32.8249
22.6041 | -35.3813 | -32.8311
22.6102 | -35.3817 | -32.8373
22.6163 | -35.3821 | -32.8432
22.6224 | -35.3825 | -32.8495
22.6285 | -35.383 | -32.8556
22.6346 | -35.3834 | -32.862
22.6407 | -35.3838 | -32.8681
22.6468 | -35.3842 | -32.874
22.6529 | -35.3846 | -32.8798
22.659 | -35.3851 | -32.886
22.6651 | -35.3855 | -32.8922
22.6712 | -35.3859 | -32.8985
22.6773 | -35.3863 | -32.9046
22.6834 | -35.3867 | -32.9109
22.6895 | -35.3871 | -32.917
22.6956 | -35.3875 | -32.9234
22.7017 | -35.3879 | -32.9296
22.7079 | -35.3883 | -32.9357
22.714 | -35.3887 | -32.9419
22.7201 | -35.3892 | -32.9479
22.7262 | -35.3895 | -32.9542
22.7323 | -35.39 | -32.9601
22.7384 | -35.3904 | -32.9663
22.7445 | -35.3908 | -32.9725
22.7506 | -35.3912 | -32.9784
22.7567 | -35.3916 | -32.9845
22.7628 | -35.392 | -32.9905
22.7689 | -35.3924 | -32.9967
22.775 | -35.3929 | -33.0031
22.7811 | -35.3933 | -33.0093
22.7872 | -35.3937 | -33.0154
22.7933 | -35.3941 | -33.0216
22.7994 | -35.3945 | -33.0278
22.8055 | -35.3949 | -33.0341
22.8116 | -35.3953 | -33.0399
22.8177 | -35.3956 | -33.046
22.8238 | -35.3961 | -33.0521
22.8299 | -35.3965 | -33.0581
22.836 | -35.397 | -33.0644
22.8421 | -35.3973 | -33.0705
22.8482 | -35.3977 | -33.0765
22.8544 | -35.3981 | -33.0827
22.8605 | -35.3986 | -33.089
22.8666 | -35.399 | -33.0948
22.8727 | -35.3994 | -33.101
22.8788 | -35.3998 | -33.1072
22.8849 | -35.4002 | -33.1133
22.891 | -35.4006 | -33.1193
22.8971 | -35.401 | -33.1255
22.9032 | -35.4014 | -33.1315
22.9093 | -35.4019 | -33.1379
22.9154 | -35.4023 | -33.1438
22.9215 | -35.4027 | -33.1499
22.9276 | -35.4031 | -33.156
22.9337 | -35.4035 | -33.1617
22.9398 | -35.4039 | -33.1676
22.9459 | -35.4043 | -33.1738
22.952 | -35.4047 | -33.1797
22.9581 | -35.4051 | -33.1858
22.9642 | -35.4055 | -33.192
22.9703 | -35.406 | -33.1979
22.9764 | -35.4064 | -33.2039
22.9825 | -35.4068 | -33.21
22.9886 | -35.4073 | -33.2161
22.9948 | -35.4077 | -33.2222
23.0009 | -35.4081 | -33.2284
23.007 | -35.4086 | -33.2343
23.0131 | -35.409 | -33.2405
23.0192 | -35.4094 | -33.2467
23.0253 | -35.4098 | -33.2528
23.0314 | -35.4102 | -33.259
23.0375 | -35.4106 | -33.2651
23.0436 | -35.4111 | -33.2708
23.0497 | -35.4115 | -33.2772
23.0558 | -35.4119 | -33.2833
23.0619 | -35.4123 | -33.2894
23.068 | -35.4127 | -33.2955
23.0741 | -35.4131 | -33.3012
23.0802 | -35.4136 | -33.3072
23.0863 | -35.414 | -33.3131
23.0924 | -35.4144 | -33.319
23.0985 | -35.4148 | -33.3251
23.1046 | -35.4152 | -33.3312
23.1107 | -35.4156 | -33.3372
23.1168 | -35.416 | -33.3432
23.1229 | -35.4165 | -33.3493
23.129 | -35.4169 | -33.3553
23.1351 | -35.4173 | -33.3612
23.1413 | -35.4177 | -33.3673
23.1474 | -35.4181 | -33.3727
23.1535 | -35.4185 | -33.3786
23.1596 | -35.4189 | -33.3844
23.1657 | -35.4193 | -33.3903
23.1718 | -35.4197 | -33.3964
23.1779 | -35.4202 | -33.4028
23.184 | -35.4206 | -33.4085
23.1901 | -35.421 | -33.4145
23.1962 | -35.4215 | -33.4206
23.2023 | -35.4219 | -33.4267
23.2084 | -35.4222 | -33.4328
23.2145 | -35.4227 | -33.4383
23.2206 | -35.4231 | -33.4442
23.2267 | -35.4235 | -33.4504
23.2328 | -35.424 | -33.4563
23.2389 | -35.4244 | -33.4622
23.245 | -35.4248 | -33.4681
23.2511 | -35.4252 | -33.4739
23.2572 | -35.4256 | -33.48
23.2633 | -35.426 | -33.486
23.2694 | -35.4265 | -33.4921
23.2755 | -35.4269 | -33.4981
23.2817 | -35.4273 | -33.5041
23.2878 | -35.4277 | -33.5099
23.2939 | -35.4281 | -33.5161
23.3 | -35.4285 | -33.5222
23.3061 | -35.4289 | -33.5282
23.3122 | -35.4294 | -33.5341
23.3183 | -35.4298 | -33.5403
23.3244 | -35.4302 | -33.5463
23.3305 | -35.4306 | -33.5524
23.3366 | -35.4311 | -33.5583
23.3427 | -35.4315 | -33.5642
23.3488 | -35.4319 | -33.5701
23.3549 | -35.4323 | -33.5761
23.361 | -35.4327 | -33.5816
23.3671 | -35.4332 | -33.5876
23.3732 | -35.4336 | -33.5936
23.3793 | -35.434 | -33.5994
23.3854 | -35.4345 | -33.6057
23.3915 | -35.4349 | -33.6114
23.3976 | -35.4353 | -33.6176
23.4037 | -35.4357 | -33.6236
23.4098 | -35.4361 | -33.6296
23.4159 | -35.4366 | -33.6356
23.422 | -35.437 | -33.6415
23.4282 | -35.4373 | -33.6474
23.4343 | -35.4377 | -33.6533
23.4404 | -35.4381 | -33.6596
23.4465 | -35.4386 | -33.6654
23.4526 | -35.4389 | -33.6714
23.4587 | -35.4393 | -33.6773
23.4648 | -35.4397 | -33.6831
23.4709 | -35.4402 | -33.6887
23.477 | -35.4406 | -33.6947
23.4831 | -35.441 | -33.7008
23.4892 | -35.4413 | -33.7068
23.4953 | -35.4418 | -33.7127
23.5014 | -35.4422 | -33.7185
23.5075 | -35.4426 | -33.7243
23.5136 | -35.443 | -33.7302
23.5197 | -35.4434 | -33.7359
23.5258 | -35.4439 | -33.7419
23.5319 | -35.4443 | -33.7478
23.538 | -35.4447 | -33.7536
23.5441 | -35.4452 | -33.7594
23.5502 | -35.4455 | -33.7651
23.5563 | -35.4459 | -33.771
23.5624 | -35.4463 | -33.7768
23.5686 | -35.4467 | -33.7824
23.5747 | -35.4472 | -33.7884
23.5808 | -35.4476 | -33.7941
23.5869 | -35.448 | -33.8001
23.593 | -35.4484 | -33.806
23.5991 | -35.4488 | -33.8118
23.6052 | -35.4493 | -33.8173
23.6113 | -35.4497 | -33.8235
23.6174 | -35.4501 | -33.8295
23.6235 | -35.4505 | -33.8353
23.6296 | -35.451 | -33.8412
23.6357 | -35.4514 | -33.8471
23.6418 | -35.4518 | -33.8529
23.6479 | -35.4523 | -33.8585
23.654 | -35.4527 | -33.8645
23.6601 | -35.4531 | -33.8702
23.6662 | -35.4535 | -33.8762
23.6723 | -35.4539 | -33.8817
23.6784 | -35.4543 | -33.8876
23.6845 | -35.4546 | -33.8933
23.6906 | -35.455 | -33.8992
23.6967 | -35.4554 | -33.9051
23.7028 | -35.4558 | -33.9109
23.7089 | -35.4562 | -33.9164
23.7151 | -35.4566 | -33.9221
23.7212 | -35.457 | -33.9278
23.7273 | -35.4574 | -33.9336
23.7334 | -35.4577 | -33.9396
23.7395 | -35.4581 | -33.9455
23.7456 | -35.4586 | -33.9512
23.7517 | -35.459 | -33.9568
23.7578 | -35.4594 | -33.9625
23.7639 | -35.4597 | -33.9684
23.77 | -35.4602 | -33.9743
23.7761 | -35.4606 | -33.9802
23.7822 | -35.4609 | -33.9862
23.7883 | -35.4613 | -33.9917
23.7944 | -35.4617 | -33.9976
23.8005 | -35.4621 | -34.0036
23.8066 | -35.4626 | -34.0096
23.8127 | -35.4629 | -34.0153
23.8188 | -35.4633 | -34.0213
23.8249 | -35.4638 | -34.0273
23.831 | -35.4642 | -34.0331
23.8371 | -35.4646 | -34.039
23.8432 | -35.465 | -34.0449
23.8493 | -35.4654 | -34.0506
23.8555 | -35.4658 | -34.0566
23.8616 | -35.4662 | -34.0624
23.8677 | -35.4666 | -34.068
23.8738 | -35.467 | -34.0738
23.8799 | -35.4674 | -34.0795
23.886 | -35.4678 | -34.0849
23.8921 | -35.4682 | -34.0908
23.8982 | -35.4686 | -34.0968
23.9043 | -35.469 | -34.1026
23.9104 | -35.4694 | -34.1083
23.9165 | -35.4698 | -34.1142
23.9226 | -35.4702 | -34.1196
23.9287 | -35.4706 | -34.1256
23.9348 | -35.4711 | -34.1315
23.9409 | -35.4715 | -34.1373
23.947 | -35.4719 | -34.143
23.9531 | -35.4723 | -34.1487
23.9592 | -35.4727 | -34.1545
23.9653 | -35.4731 | -34.1602
23.9714 | -35.4735 | -34.1661
23.9775 | -35.4739 | -34.1719
23.9836 | -35.4743 | -34.1778
23.9897 | -35.4748 | -34.1832
23.9958 | -35.4752 | -34.1889
24.002 | -35.4756 | -34.1945
24.0081 | -35.476 | -34.2004
24.0142 | -35.4764 | -34.2061
24.0203 | -35.4768 | -34.212
24.0264 | -35.4772 | -34.2175
24.0325 | -35.4776 | -34.2235
24.0386 | -35.478 | -34.2294
24.0447 | -35.4784 | -34.235
24.0508 | -35.4788 | -34.2406
24.0569 | -35.4792 | -34.2464
24.063 | -35.4796 | -34.2518
24.0691 | -35.48 | -34.2575
24.0752 | -35.4804 | -34.2634
24.0813 | -35.4808 | -34.2691
24.0874 | -35.4812 | -34.275
24.0935 | -35.4816 | -34.2808
24.0996 | -35.4821 | -34.2864
24.1057 | -35.4825 | -34.2922
24.1118 | -35.4829 | -34.298
24.1179 | -35.4833 | -34.3037
24.124 | -35.4837 | -34.3094
24.1301 | -35.4841 | -34.3149
24.1362 | -35.4845 | -34.3205
24.1424 | -35.4849 | -34.3263
24.1485 | -35.4853 | -34.3322
24.1546 | -35.4857 | -34.3379
24.1607 | -35.4861 | -34.3434
24.1668 | -35.4864 | -34.3489
24.1729 | -35.4868 | -34.3546
24.179 | -35.4872 | -34.3605
24.1851 | -35.4876 | -34.3664
24.1912 | -35.488 | -34.3718
24.1973 | -35.4884 | -34.3777
24.2034 | -35.4888 | -34.3833
24.2095 | -35.4892 | -34.3891
24.2156 | -35.4896 | -34.3949
24.2217 | -35.4901 | -34.4006
24.2278 | -35.4904 | -34.4059
24.2339 | -35.4908 | -34.4116
24.24 | -35.4912 | -34.4167
24.2461 | -35.4916 | -34.4223
24.2522 | -35.492 | -34.4282
24.2583 | -35.4924 | -34.4338
24.2644 | -35.4929 | -34.4395
24.2705 | -35.4933 | -34.4451
24.2766 | -35.4937 | -34.4507
24.2827 | -35.4941 | -34.4564
24.2889 | -35.4945 | -34.4623
24.295 | -35.4948 | -34.468
24.3011 | -35.4952 | -34.4736
24.3072 | -35.4956 | -34.4789
24.3133 | -35.496 | -34.4844
24.3194 | -35.4964 | -34.4903
24.3255 | -35.4968 | -34.4961
24.3316 | -35.4973 | -34.5016
24.3377 | -35.4976 | -34.5073
24.3438 | -35.498 | -34.5128
24.3499 | -35.4984 | -34.5183
24.356 | -35.4988 | -34.5239
24.3621 | -35.4992 | -34.5294
24.3682 | -35.4996 | -34.535
24.3743 | -35.5 | -34.5408
24.3804 | -35.5004 | -34.5463
24.3865 | -35.5008 | -34.5518
24.3926 | -35.5012 | -34.5575
24.3987 | -35.5016 | -34.5633
24.4048 | -35.502 | -34.569
24.4109 | -35.5023 | -34.5746
24.417 | -35.5028 | -34.5802
24.4231 | -35.5032 | -34.5859
24.4293 | -35.5036 | -34.5913
24.4354 | -35.5039 | -34.5968
24.4415 | -35.5044 | -34.6023
24.4476 | -35.5047 | -34.6076
24.4537 | -35.5052 | -34.6132
24.4598 | -35.5056 | -34.6186
24.4659 | -35.506 | -34.6243
24.472 | -35.5064 | -34.6299
24.4781 | -35.5068 | -34.6356
24.4842 | -35.5072 | -34.6409
24.4903 | -35.5076 | -34.6466
24.4964 | -35.508 | -34.6519
24.5025 | -35.5084 | -34.6575
24.5086 | -35.5088 | -34.6632
24.5147 | -35.5093 | -34.6689
24.5208 | -35.5096 | -34.6743
24.5269 | -35.51 | -34.6798
24.533 | -35.5104 | -34.6855
24.5391 | -35.5108 | -34.6911
24.5452 | -35.5112 | -34.6967
24.5513 | -35.5116 | -34.7019
24.5574 | -35.512 | -34.7075
24.5635 | -35.5124 | -34.7131
24.5696 | -35.5127 | -34.7188
24.5758 | -35.5131 | -34.7244
24.5819 | -35.5136 | -34.7298
24.588 | -35.5139 | -34.7353
24.5941 | -35.5143 | -34.7407
24.6002 | -35.5147 | -34.7464
24.6063 | -35.5151 | -34.7521
24.6124 | -35.5155 | -34.7576
24.6185 | -35.516 | -34.7631
24.6246 | -35.5163 | -34.7686
24.6307 | -35.5167 | -34.7742
24.6368 | -35.5171 | -34.7799
24.6429 | -35.5175 | -34.7853
24.649 | -35.518 | -34.7911
24.6551 | -35.5184 | -34.7966
24.6612 | -35.5188 | -34.802
24.6673 | -35.5192 | -34.8078
24.6734 | -35.5195 | -34.8136
24.6795 | -35.5199 | -34.8193
24.6856 | -35.5203 | -34.825
24.6917 | -35.5208 | -34.8307
24.6978 | -35.5211 | -34.836
24.7039 | -35.5215 | -34.8416
24.71 | -35.5219 | -34.8473
24.7162 | -35.5223 | -34.8531
24.7223 | -35.5227 | -34.8586
24.7284 | -35.5231 | -34.8639
24.7345 | -35.5235 | -34.8691
24.7406 | -35.5238 | -34.8747
24.7467 | -35.5242 | -34.8802
24.7528 | -35.5247 | -34.8857
24.7589 | -35.5251 | -34.891
24.765 | -35.5254 | -34.8962
24.7711 | -35.5258 | -34.9015
24.7772 | -35.5262 | -34.9069
24.7833 | -35.5266 | -34.9125
24.7894 | -35.527 | -34.9179
24.7955 | -35.5274 | -34.9237
24.8016 | -35.5278 | -34.929
24.8077 | -35.5281 | -34.9346
24.8138 | -35.5285 | -34.9403
24.8199 | -35.5289 | -34.9458
24.826 | -35.5293 | -34.9513
24.8321 | -35.5297 | -34.9569
24.8382 | -35.5301 | -34.9623
24.8443 | -35.5305 | -34.9676
24.8504 | -35.5309 | -34.9733
24.8565 | -35.5313 | -34.9789
24.8627 | -35.5316 | -34.9845
24.8688 | -35.532 | -34.9898
24.8749 | -35.5324 | -34.9955
24.881 | -35.5328 | -35.001
24.8871 | -35.5332 | -35.0066
24.8932 | -35.5336 | -35.0121
24.8993 | -35.5339 | -35.0175
24.9054 | -35.5343 | -35.023
24.9115 | -35.5347 | -35.0285
24.9176 | -35.5351 | -35.0339
24.9237 | -35.5355 | -35.0394
24.9298 | -35.5359 | -35.045
24.9359 | -35.5363 | -35.0506
24.942 | -35.5367 | -35.0559
24.9481 | -35.5371 | -35.0616
24.9542 | -35.5375 | -35.0671
24.9603 | -35.5379 | -35.0727
24.9664 | -35.5383 | -35.0783
24.9725 | -35.5387 | -35.0836
24.9786 | -35.539 | -35.0889
24.9847 | -35.5395 | -35.0944
24.9908 | -35.5398 | -35.0998
24.9969 | -35.5402 | -35.1053
25.0031 | -35.5406 | -35.1108
25.0092 | -35.541 | -35.1159
25.0153 | -35.5414 | -35.1213
25.0214 | -35.5418 | -35.1267
25.0275 | -35.5422 | -35.1321
25.0336 | -35.5426 | -35.1375
25.0397 | -35.543 | -35.1432
25.0458 | -35.5433 | -35.1485
25.0519 | -35.5437 | -35.154
25.058 | -35.5441 | -35.1594
25.0641 | -35.5445 | -35.1648
25.0702 | -35.5449 | -35.1702
25.0763 | -35.5453 | -35.1757
25.0824 | -35.5456 | -35.1808
25.0885 | -35.5461 | -35.1862
25.0946 | -35.5465 | -35.1916
25.1007 | -35.5468 | -35.1971
25.1068 | -35.5472 | -35.2027
25.1129 | -35.5476 | -35.2081
25.119 | -35.548 | -35.2134
25.1251 | -35.5484 | -35.2189
25.1312 | -35.5487 | -35.2243
25.1373 | -35.5491 | -35.23
25.1435 | -35.5495 | -35.2356
25.1496 | -35.5499 | -35.2407
25.1557 | -35.5503 | -35.2461
25.1618 | -35.5507 | -35.2514
25.1679 | -35.5511 | -35.257
25.174 | -35.5514 | -35.2626
25.1801 | -35.5518 | -35.2681
25.1862 | -35.5522 | -35.2735
25.1923 | -35.5525 | -35.2791
25.1984 | -35.5529 | -35.2848
25.2045 | -35.5533 | -35.2904
25.2106 | -35.5537 | -35.2958
25.2167 | -35.554 | -35.3013
25.2228 | -35.5543 | -35.307
25.2289 | -35.5547 | -35.3124
25.235 | -35.5551 | -35.318
25.2411 | -35.5555 | -35.3235
25.2472 | -35.5558 | -35.3291
25.2533 | -35.5562 | -35.3346
25.2594 | -35.5565 | -35.3398
25.2655 | -35.5569 | -35.3453
25.2716 | -35.5572 | -35.3507
25.2777 | -35.5576 | -35.3561
25.2838 | -35.558 | -35.3617
25.29 | -35.5584 | -35.3669
25.2961 | -35.5588 | -35.3721
25.3022 | -35.5592 | -35.3775
25.3083 | -35.5596 | -35.3832
25.3144 | -35.56 | -35.3885
25.3205 | -35.5603 | -35.394
25.3266 | -35.5607 | -35.3992
25.3327 | -35.5611 | -35.4047
25.3388 | -35.5615 | -35.4101
25.3449 | -35.5619 | -35.4155
25.351 | -35.5623 | -35.421
25.3571 | -35.5626 | -35.4264
25.3632 | -35.563 | -35.4317
25.3693 | -35.5634 | -35.4371
25.3754 | -35.5637 | -35.4425
25.3815 | -35.5641 | -35.4478
25.3876 | -35.5645 | -35.4532
25.3937 | -35.5648 | -35.4586
25.3998 | -35.5652 | -35.4636
25.4059 | -35.5656 | -35.4692
25.412 | -35.566 | -35.4747
25.4181 | -35.5664 | -35.4802
25.4242 | -35.5667 | -35.4856
25.4304 | -35.5671 | -35.4912
25.4365 | -35.5675 | -35.4964
25.4426 | -35.5679 | -35.502
25.4487 | -35.5682 | -35.5073
25.4548 | -35.5686 | -35.5128
25.4609 | -35.569 | -35.5183
25.467 | -35.5695 | -35.5233
25.4731 | -35.5698 | -35.5289
25.4792 | -35.5702 | -35.5344
25.4853 | -35.5706 | -35.5398
25.4914 | -35.571 | -35.5453
25.4975 | -35.5713 | -35.5507
25.5036 | -35.5716 | -35.5562
25.5097 | -35.572 | -35.5615
25.5158 | -35.5724 | -35.5668
25.5219 | -35.5728 | -35.5722
25.528 | -35.5731 | -35.5776
25.5341 | -35.5735 | -35.5829
25.5402 | -35.5739 | -35.588
25.5463 | -35.5743 | -35.5933
25.5524 | -35.5747 | -35.5988
25.5585 | -35.575 | -35.6045
25.5646 | -35.5754 | -35.6099
25.5707 | -35.5758 | -35.6153
25.5769 | -35.5762 | -35.6202
25.583 | -35.5766 | -35.6255
25.5891 | -35.5769 | -35.631
25.5952 | -35.5773 | -35.6363
25.6013 | -35.5777 | -35.6417
25.6074 | -35.5781 | -35.647
25.6135 | -35.5785 | -35.6528
25.6196 | -35.5789 | -35.6582
25.6257 | -35.5793 | -35.6636
25.6318 | -35.5796 | -35.6689
25.6379 | -35.58 | -35.6744
25.644 | -35.5803 | -35.6798
25.6501 | -35.5807 | -35.6851
25.6562 | -35.5811 | -35.6905
25.6623 | -35.5815 | -35.6959
25.6684 | -35.5819 | -35.7013
25.6745 | -35.5823 | -35.7065
25.6806 | -35.5826 | -35.712
25.6867 | -35.583 | -35.7175
25.6928 | -35.5834 | -35.7232
25.6989 | -35.5838 | -35.7288
25.705 | -35.5841 | -35.734
25.7111 | -35.5845 | -35.7395
25.7173 | -35.5849 | -35.7449
25.7234 | -35.5852 | -35.7501
25.7295 | -35.5856 | -35.7553
25.7356 | -35.586 | -35.7607
25.7417 | -35.5863 | -35.766
25.7478 | -35.5866 | -35.7713
25.7539 | -35.587 | -35.7766
25.76 | -35.5873 | -35.7819
25.7661 | -35.5877 | -35.7872
25.7722 | -35.5881 | -35.7927
25.7783 | -35.5884 | -35.7981
25.7844 | -35.5888 | -35.8034
25.7905 | -35.5892 | -35.8087
25.7966 | -35.5896 | -35.814
25.8027 | -35.5899 | -35.8194
25.8088 | -35.5903 | -35.8251
25.8149 | -35.5906 | -35.8307
25.821 | -35.5909 | -35.8361
25.8271 | -35.5913 | -35.8416
25.8332 | -35.5917 | -35.8471
25.8393 | -35.5921 | -35.8526
25.8454 | -35.5925 | -35.8581
25.8515 | -35.5928 | -35.8634
25.8576 | -35.5932 | -35.8688
25.8638 | -35.5936 | -35.8742
25.8699 | -35.5939 | -35.8797
25.876 | -35.5943 | -35.885
25.8821 | -35.5946 | -35.8904
25.8882 | -35.5949 | -35.8959
25.8943 | -35.5953 | -35.9016
25.9004 | -35.5956 | -35.907
25.9065 | -35.596 | -35.9128
25.9126 | -35.5963 | -35.9182
25.9187 | -35.5967 | -35.9232
25.9248 | -35.5971 | -35.9284
25.9309 | -35.5975 | -35.9339
25.937 | -35.5978 | -35.9395
25.9431 | -35.5982 | -35.9447
25.9492 | -35.5985 | -35.9496
25.9553 | -35.5989 | -35.9547
25.9614 | -35.5992 | -35.9601
25.9675 | -35.5996 | -35.9652
25.9736 | -35.6 | -35.9708
25.9797 | -35.6003 | -35.9763
25.9858 | -35.6007 | -35.9816
25.9919 | -35.601 | -35.9869
25.998 | -35.6014 | -35.9921
26.0042 | -35.6017 | -35.9974
26.0103 | -35.6021 | -36.0028
26.0164 | -35.6024 | -36.0082
26.0225 | -35.6028 | -36.0138
26.0286 | -35.6032 | -36.019
26.0347 | -35.6035 | -36.0242
26.0408 | -35.6038 | -36.0296
26.0469 | -35.6042 | -36.035
26.053 | -35.6045 | -36.0405
26.0591 | -35.6049 | -36.0458
26.0652 | -35.6052 | -36.0508
26.0713 | -35.6055 | -36.0563
26.0774 | -35.6059 | -36.0615
26.0835 | -35.6063 | -36.0671
26.0896 | -35.6066 | -36.0726
26.0957 | -35.607 | -36.078
26.1018 | -35.6074 | -36.0832
26.1079 | -35.6077 | -36.0887
26.114 | -35.6081 | -36.0939
26.1201 | -35.6084 | -36.0993
26.1262 | -35.6088 | -36.1047
26.1323 | -35.6091 | -36.11
26.1384 | -35.6095 | -36.1152
26.1445 | -35.6098 | -36.1207
26.1507 | -35.6101 | -36.126
26.1568 | -35.6105 | -36.1312
26.1629 | -35.6108 | -36.1364
26.169 | -35.6111 | -36.1416
26.1751 | -35.6115 | -36.1467
26.1812 | -35.6119 | -36.1521
26.1873 | -35.6122 | -36.1574
26.1934 | -35.6125 | -36.163
26.1995 | -35.6128 | -36.1685
26.2056 | -35.6132 | -36.1736
26.2117 | -35.6135 | -36.1789
26.2178 | -35.6138 | -36.1844
26.2239 | -35.6142 | -36.1895
26.23 | -35.6145 | -36.1949
26.2361 | -35.6149 | -36.2
26.2422 | -35.6152 | -36.205
26.2483 | -35.6156 | -36.2106
26.2544 | -35.6159 | -36.2158
26.2605 | -35.6163 | -36.221
26.2666 | -35.6166 | -36.2262
26.2727 | -35.617 | -36.2314
26.2788 | -35.6173 | -36.2366
26.2849 | -35.6177 | -36.2419
26.2911 | -35.6181 | -36.2472
26.2972 | -35.6184 | -36.2525
26.3033 | -35.6188 | -36.2577
26.3094 | -35.6191 | -36.2628
26.3155 | -35.6194 | -36.268
26.3216 | -35.6198 | -36.2734
26.3277 | -35.6201 | -36.2789
26.3338 | -35.6205 | -36.2842
26.3399 | -35.6208 | -36.2894
26.346 | -35.6212 | -36.2947
26.3521 | -35.6215 | -36.3003
26.3582 | -35.6219 | -36.3056
26.3643 | -35.6222 | -36.311
26.3704 | -35.6225 | -36.3164
26.3765 | -35.6228 | -36.3217
26.3826 | -35.6231 | -36.3267
26.3887 | -35.6235 | -36.332
26.3948 | -35.6239 | -36.3369
26.4009 | -35.6242 | -36.3424
26.407 | -35.6246 | -36.3478
26.4131 | -35.625 | -36.3535
26.4192 | -35.6253 | -36.3584
26.4253 | -35.6256 | -36.364
26.4314 | -35.626 | -36.3691
26.4376 | -35.6264 | -36.3742
26.4437 | -35.6267 | -36.3795
26.4498 | -35.6271 | -36.3846
26.4559 | -35.6274 | -36.3898
26.462 | -35.6277 | -36.3951
26.4681 | -35.6281 | -36.4002
26.4742 | -35.6285 | -36.4055
26.4803 | -35.6288 | -36.4108
26.4864 | -35.6291 | -36.4157
26.4925 | -35.6294 | -36.4212
26.4986 | -35.6298 | -36.4263
26.5047 | -35.6301 | -36.4313
26.5108 | -35.6305 | -36.4366
26.5169 | -35.6308 | -36.4417
26.523 | -35.6311 | -36.4469
26.5291 | -35.6315 | -36.4519
26.5352 | -35.6318 | -36.4572
26.5413 | -35.6321 | -36.4625
26.5474 | -35.6325 | -36.4678
26.5535 | -35.6329 | -36.4733
26.5596 | -35.6332 | -36.4786
26.5657 | -35.6336 | -36.4839
26.5718 | -35.6339 | -36.489
26.578 | -35.6342 | -36.4943
26.5841 | -35.6345 | -36.4992
26.5902 | -35.6348 | -36.5044
26.5963 | -35.6352 | -36.5095
26.6024 | -35.6355 | -36.515
26.6085 | -35.6359 | -36.5203
26.6146 | -35.6362 | -36.526
26.6207 | -35.6365 | -36.5313
26.6268 | -35.6369 | -36.5368
26.6329 | -35.6371 | -36.5417
26.639 | -35.6374 | -36.5468
26.6451 | -35.6378 | -36.552
26.6512 | -35.6382 | -36.5574
26.6573 | -35.6385 | -36.5627
26.6634 | -35.6389 | -36.5677
26.6695 | -35.6392 | -36.5733
26.6756 | -35.6395 | -36.5787
26.6817 | -35.6399 | -36.584
26.6878 | -35.6402 | -36.5895
26.6939 | -35.6405 | -36.5944
26.7 | -35.6408 | -36.5994
26.7061 | -35.6411 | -36.6044
26.7122 | -35.6415 | -36.6096
26.7183 | -35.6419 | -36.6149
26.7245 | -35.6422 | -36.6199
26.7306 | -35.6426 | -36.6249
26.7367 | -35.6429 | -36.63
26.7428 | -35.6432 | -36.6354
26.7489 | -35.6436 | -36.641
26.755 | -35.644 | -36.6463
26.7611 | -35.6443 | -36.6518
26.7672 | -35.6446 | -36.6573
26.7733 | -35.6449 | -36.6625
26.7794 | -35.6453 | -36.668
26.7855 | -35.6456 | -36.6732
26.7916 | -35.6459 | -36.6783
26.7977 | -35.6463 | -36.6836
26.8038 | -35.6466 | -36.6892
26.8099 | -35.647 | -36.6946
26.816 | -35.6473 | -36.7001
26.8221 | -35.6476 | -36.7051
26.8282 | -35.648 | -36.7104
26.8343 | -35.6484 | -36.7159
26.8404 | -35.6487 | -36.7213
26.8465 | -35.649 | -36.7268
26.8526 | -35.6493 | -36.7322
26.8587 | -35.6496 | -36.7372
26.8649 | -35.65 | -36.7422
26.871 | -35.6503 | -36.7474
26.8771 | -35.6506 | -36.7522
26.8832 | -35.6509 | -36.7575
26.8893 | -35.6513 | -36.7625
26.8954 | -35.6516 | -36.7677
26.9015 | -35.6519 | -36.7731
26.9076 | -35.6522 | -36.7784
26.9137 | -35.6525 | -36.783
26.9198 | -35.6528 | -36.7885
26.9259 | -35.6531 | -36.7934
26.932 | -35.6534 | -36.7987
26.9381 | -35.6537 | -36.8038
26.9442 | -35.654 | -36.8087
26.9503 | -35.6544 | -36.814
26.9564 | -35.6546 | -36.8191
26.9625 | -35.655 | -36.8243
26.9686 | -35.6554 | -36.8294
26.9747 | -35.6558 | -36.8348
26.9808 | -35.6561 | -36.8398
26.9869 | -35.6564 | -36.8452
26.993 | -35.6567 | -36.8505
26.9991 | -35.657 | -36.8558
27.0052 | -35.6574 | -36.8614
27.0114 | -35.6577 | -36.8665
27.0175 | -35.658 | -36.8713
27.0236 | -35.6584 | -36.8766
27.0297 | -35.6587 | -36.8815
27.0358 | -35.659 | -36.8867
27.0419 | -35.6594 | -36.8917
27.048 | -35.6597 | -36.8973
27.0541 | -35.66 | -36.9021
27.0602 | -35.6603 | -36.907
27.0663 | -35.6607 | -36.912
27.0724 | -35.661 | -36.9171
27.0785 | -35.6614 | -36.9225
27.0846 | -35.6617 | -36.9273
27.0907 | -35.662 | -36.9326
27.0968 | -35.6623 | -36.938
27.1029 | -35.6626 | -36.9433
27.109 | -35.6629 | -36.9479
27.1151 | -35.6632 | -36.9532
27.1212 | -35.6635 | -36.9582
27.1273 | -35.6639 | -36.9632
27.1334 | -35.6642 | -36.9687
27.1395 | -35.6646 | -36.974
27.1456 | -35.6649 | -36.979
27.1518 | -35.6652 | -36.9846
27.1579 | -35.6655 | -36.9896
27.164 | -35.6659 | -36.9948
27.1701 | -35.6662 | -37
27.1762 | -35.6665 | -37.0053
27.1823 | -35.6668 | -37.0107
27.1884 | -35.6671 | -37.016
27.1945 | -35.6676 | -37.0208
27.2006 | -35.6679 | -37.0257
27.2067 | -35.6682 | -37.0306
27.2128 | -35.6685 | -37.0361
27.2189 | -35.6688 | -37.0414
27.225 | -35.6691 | -37.047
27.2311 | -35.6694 | -37.052
27.2372 | -35.6697 | -37.0574
27.2433 | -35.6701 | -37.0625
27.2494 | -35.6704 | -37.0679
27.2555 | -35.6707 | -37.0731
27.2616 | -35.6711 | -37.0783
27.2677 | -35.6714 | -37.0839
27.2738 | -35.6718 | -37.0888
27.2799 | -35.6721 | -37.0943
27.286 | -35.6724 | -37.0991
27.2921 | -35.6727 | -37.1047
27.2983 | -35.673 | -37.1092
27.3044 | -35.6733 | -37.1147
27.3105 | -35.6736 | -37.12
27.3166 | -35.674 | -37.1253
27.3227 | -35.6743 | -37.13
27.3288 | -35.6746 | -37.1347
27.3349 | -35.6749 | -37.1392
27.341 | -35.6752 | -37.1444
27.3471 | -35.6756 | -37.1496
27.3532 | -35.6759 | -37.1548
27.3593 | -35.6763 | -37.1603
27.3654 | -35.6766 | -37.1656
27.3715 | -35.6769 | -37.1704
27.3776 | -35.6772 | -37.1758
27.3837 | -35.6776 | -37.1809
27.3898 | -35.6779 | -37.1857
27.3959 | -35.6782 | -37.1908
27.402 | -35.6784 | -37.1958
27.4081 | -35.6788 | -37.2012
27.4142 | -35.679 | -37.2065
27.4203 | -35.6793 | -37.2114
27.4264 | -35.6796 | -37.2169
27.4325 | -35.6799 | -37.2218
27.4387 | -35.6802 | -37.2265
27.4448 | -35.6805 | -37.2315
27.4509 | -35.6807 | -37.2368
27.457 | -35.6811 | -37.2417
27.4631 | -35.6813 | -37.2471
27.4692 | -35.6817 | -37.2523
27.4753 | -35.682 | -37.257
27.4814 | -35.6823 | -37.2623
27.4875 | -35.6825 | -37.2673
27.4936 | -35.6828 | -37.2729
27.4997 | -35.6832 | -37.2779
27.5058 | -35.6835 | -37.283
27.5119 | -35.6838 | -37.288
27.518 | -35.6841 | -37.2933
27.5241 | -35.6844 | -37.2982
27.5302 | -35.6847 | -37.3037
27.5363 | -35.685 | -37.3089
27.5424 | -35.6853 | -37.3143
27.5485 | -35.6856 | -37.3197
27.5546 | -35.686 | -37.325
27.5607 | -35.6863 | -37.33
27.5668 | -35.6865 | -37.3352
27.5729 | -35.6869 | -37.3401
27.5791 | -35.6872 | -37.3452
27.5852 | -35.6875 | -37.3499
27.5913 | -35.6878 | -37.3546
27.5974 | -35.6881 | -37.3593
27.6035 | -35.6884 | -37.3648
27.6096 | -35.6887 | -37.3698
27.6157 | -35.689 | -37.3745
27.6218 | -35.6893 | -37.3798
27.6279 | -35.6896 | -37.3853
27.634 | -35.6899 | -37.3907
27.6401 | -35.6903 | -37.3964
27.6462 | -35.6906 | -37.4015
27.6523 | -35.6908 | -37.4062
27.6584 | -35.6911 | -37.4115
27.6645 | -35.6914 | -37.4168
27.6706 | -35.6918 | -37.4218
27.6767 | -35.6921 | -37.4268
27.6828 | -35.6923 | -37.4321
27.6889 | -35.6927 | -37.4374
27.695 | -35.693 | -37.4428
27.7011 | -35.6933 | -37.4479
27.7072 | -35.6936 | -37.4526
27.7133 | -35.694 | -37.4577
27.7194 | -35.6942 | -37.4622
27.7256 | -35.6945 | -37.4674
27.7317 | -35.6949 | -37.4723
27.7378 | -35.6952 | -37.477
27.7439 | -35.6955 | -37.4822
27.75 | -35.6958 | -37.4873
27.7561 | -35.6961 | -37.4925
27.7622 | -35.6963 | -37.4975
27.7683 | -35.6966 | -37.5024
27.7744 | -35.697 | -37.5072
27.7805 | -35.6973 | -37.5122
27.7866 | -35.6976 | -37.5173
27.7927 | -35.6978 | -37.5215
27.7988 | -35.6981 | -37.5264
27.8049 | -35.6984 | -37.5315
27.811 | -35.6987 | -37.5368
27.8171 | -35.699 | -37.5419
27.8232 | -35.6993 | -37.5467
27.8293 | -35.6996 | -37.5516
27.8354 | -35.6999 | -37.5568
27.8415 | -35.7002 | -37.5615
27.8476 | -35.7005 | -37.5665
27.8537 | -35.7008 | -37.5718
27.8598 | -35.7011 | -37.5764
27.866 | -35.7014 | -37.5816
27.8721 | -35.7017 | -37.5872
27.8782 | -35.702 | -37.5923
27.8843 | -35.7024 | -37.5972
27.8904 | -35.7027 | -37.6014
27.8965 | -35.703 | -37.6065
27.9026 | -35.7033 | -37.6116
27.9087 | -35.7036 | -37.6167
27.9148 | -35.7038 | -37.6215
27.9209 | -35.7041 | -37.6267
27.927 | -35.7045 | -37.6316
27.9331 | -35.7048 | -37.6369
27.9392 | -35.7052 | -37.642
27.9453 | -35.7055 | -37.6469
27.9514 | -35.7058 | -37.6521
27.9575 | -35.706 | -37.6573
27.9636 | -35.7063 | -37.6627
27.9697 | -35.7066 | -37.6682
27.9758 | -35.7069 | -37.6731
27.9819 | -35.7072 | -37.6788
27.988 | -35.7075 | -37.684
27.9941 | -35.7077 | -37.6891
28.0002 | -35.708 | -37.694
28.0063 | -35.7083 | -37.6989
28.0125 | -35.7087 | -37.7037
28.0186 | -35.7089 | -37.7088
28.0247 | -35.7092 | -37.7139
28.0308 | -35.7095 | -37.7185
28.0369 | -35.7098 | -37.7233
28.043 | -35.7101 | -37.7283
28.0491 | -35.7104 | -37.7333
28.0552 | -35.7107 | -37.7387
28.0613 | -35.7109 | -37.7436
28.0674 | -35.7112 | -37.7488
28.0735 | -35.7116 | -37.7533

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4093 | -41.766
21.4908 | -35.6142 | -34.8965
16.4948 | -34.9781 | -27.6101
12.6602 | -34.5992 | -21.3882
9.7171 | -34.3786 | -16.5358
7.4582 | -34.2184 | -12.5848
5.7244 | -34.1067 | -9.3785
4.3936 | -34.04 | -6.7822
3.3722 | -34.017 | -4.7366
2.5883 | -33.9963 | -3.1833
1.9866 | -33.9865 | -1.9377
1.5248 | -33.9888 | -0.93859
1.1703 | -33.9888 | -0.28301
1 | -33.9862 | 0.041667
0.89822 | -33.9756 | 0.31587
0.68941 | -33.9603 | 0.90372
0.52919 | -33.951 | 1.3606
0.40612 | -33.9548 | 1.9761
0.31171 | -33.9564 | 2.4384
0.23928 | -33.9677 | 3.1665
0.18358 | -33.9772 | 3.589
0.14091 | -33.9762 | 4.4781
0.10819 | -33.9871 | 5.6471
0.083017 | -34.0121 | 7.1262
0.063705 | -34.0589 | 9.0012
0.048923 | -34.1321 | 11.4919
0.037479 | -34.2293 | 14.6846
0.028801 | -34.4017 | 18.7119
0.022078 | -34.7139 | 24.1181
0.017023 | -35.1957 | 30.5452
0.013018 | -36.0002 | 38.2228
0.010014 | -37.1276 | 46.1504

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003719 | -3.6246
0.007438 | -27.6419
0.011158 | -34.6049
0.014877 | -39.694
0.018596 | -43.1138
0.022315 | -44.7974
0.026034 | -47.6212
0.029754 | -49.5552
0.033473 | -50.9746
0.037192 | -51.6801
0.040911 | -53.0566
0.04463 | -54.5421
0.04835 | -56.0224
0.052069 | -55.8959
0.055788 | -56.6547
0.059507 | -57.5907
0.063226 | -57.5455
0.066946 | -58.7732
0.070665 | -59.9688
0.074384 | -61.3662
0.078103 | -61.585
0.081822 | -61.7508
0.085542 | -63.5305
0.089261 | -62.3694
0.09298 | -62.3639
0.096699 | -64.1449
0.10042 | -64.3712
0.10414 | -64.3437
0.10786 | -64.9835
0.11158 | -65.4083
0.11529 | -66.6523
0.11901 | -65.3294
0.12273 | -64.754
0.12645 | -66.7441
0.13017 | -66.8997
0.13389 | -65.8529
0.13761 | -66.6951
0.14133 | -67.4757
0.14505 | -67.463
0.14877 | -67.7372
0.15249 | -67.7563
0.15621 | -67.4992
0.15993 | -68.2608
0.16365 | -69.2192
0.16736 | -70.3285
0.17108 | -70.0705
0.1748 | -69.5126
0.17852 | -70.5991
0.18224 | -71.2664
0.18596 | -70.9748
0.18968 | -71.1559
0.1934 | -71.3398
0.19712 | -70.7189
0.20084 | -70.1671
0.20456 | -70.9156
0.20827 | -72.0579
0.21199 | -72.2823
0.21571 | -72.427
0.21943 | -73.0131
0.22315 | -72.7092
0.22687 | -72.3187
0.23059 | -72.0236
0.23431 | -72.0311
0.23803 | -72.4443
0.24175 | -73.2655
0.24547 | -73.2101
0.24919 | -72.4819
0.25291 | -72.909
0.25662 | -74.1446
0.26034 | -73.787
0.26406 | -73.8175
0.26778 | -74.5616
0.2715 | -75.6935
0.27522 | -76.068
0.27894 | -76.5485
0.28266 | -75.9393
0.28638 | -75.0828
0.2901 | -75.369
0.29382 | -75.0036
0.29754 | -74.4363
0.30125 | -74.5567
0.30497 | -75.9068
0.30869 | -75.7821
0.31241 | -76.2058
0.31613 | -76.5383
0.31985 | -76.7851
0.32357 | -77.4439
0.32729 | -77.5389
0.33101 | -77.7222
0.33473 | -76.7357
0.33845 | -76.311
0.34217 | -76.7828
0.34589 | -76.2385
0.3496 | -77.1892
0.35332 | -77.36
0.35704 | -77.4683
0.36076 | -77.1138
0.36448 | -77.9713
0.3682 | -78.6316
0.37192 | -77.681
0.37564 | -77.3431
0.37936 | -77.8879
0.38308 | -76.911
0.3868 | -76.4869
0.39052 | -77.3738
0.39424 | -77.5554
0.39795 | -78.2755
0.40167 | -78.6975
0.40539 | -79.6977
0.40911 | -78.9689
0.41283 | -78.5844
0.41655 | -79.3248
0.42027 | -79.9408
0.42399 | -79.5204
0.42771 | -78.6538
0.43143 | -79.2164
0.43515 | -79.7495
0.43887 | -79.4193
0.44259 | -78.5412
0.4463 | -78.7676
0.45002 | -78.8669
0.45374 | -78.8879
0.45746 | -78.7174
0.46118 | -79.1668
0.4649 | -78.9622
0.46862 | -79.9275
0.47234 | -79.959
0.47606 | -79.2528
0.47978 | -79.9843
0.4835 | -79.8517
0.48722 | -79.8478
0.49093 | -79.6944
0.49465 | -79.9675
0.49837 | -81.2992
0.50209 | -81.4591
0.50581 | -81.4411
0.50953 | -81.145
0.51325 | -80.554
0.51697 | -81.7971
0.52069 | -81.457
0.52441 | -80.6472
0.52813 | -81.1024
0.53185 | -82.2086
0.53556 | -81.9563
0.53928 | -81.0923
0.543 | -80.4846
0.54672 | -80.7978
0.55044 | -81.7107
0.55416 | -82.2451
0.55788 | -82.2034
0.5616 | -82.0628
0.56532 | -81.9999
0.56904 | -81.593
0.57276 | -81.1354
0.57648 | -80.755
0.5802 | -80.5437
0.58391 | -80.8495
0.58763 | -81.4983
0.59135 | -81.9521
0.59507 | -81.7257
0.59879 | -81.5018
0.60251 | -81.9695
0.60623 | -82.7471
0.60995 | -83.2147
0.61367 | -83.2212
0.61739 | -82.8054
0.62111 | -82.5585
0.62483 | -82.6837
0.62855 | -82.5791
0.63226 | -82.3706
0.63598 | -82.0969
0.6397 | -81.9833
0.64342 | -82.177
0.64714 | -82.5314
0.65086 | -82.9135
0.65458 | -83.0366
0.6583 | -82.7902
0.66202 | -82.3785
0.66574 | -82.062
0.66946 | -82.4095
0.67317 | -83.0606
0.67689 | -83.72
0.68061 | -83.7738
0.68433 | -83.1618
0.68805 | -83.037
0.69177 | -83.6509
0.69549 | -83.9916
0.69921 | -83.7911
0.70293 | -83.1825
0.70665 | -83.2815
0.71037 | -84.1141
0.71409 | -85.0087
0.71781 | -85.2524
0.72152 | -84.8533
0.72524 | -84.4899
0.72896 | -84.2341
0.73268 | -83.971
0.7364 | -83.7936
0.74012 | -83.6508
0.74384 | -83.6562
0.74756 | -83.8257
0.75128 | -83.9245
0.755 | -84.1491
0.75872 | -84.6894
0.76244 | -84.5789
0.76616 | -84.0038
0.76987 | -83.9967
0.77359 | -84.5755
0.77731 | -84.7688
0.78103 | -84.4362
0.78475 | -84.5937
0.78847 | -85.2478
0.79219 | -85.5819
0.79591 | -85.3636
0.79963 | -85.1168
0.80335 | -85.1035
0.80707 | -84.9201
0.81079 | -84.3859
0.81451 | -84.015
0.81822 | -84.0928
0.82194 | -84.2594
0.82566 | -84.1715
0.82938 | -84.068
0.8331 | -84.1465
0.83682 | -84.2829
0.84054 | -84.4572
0.84426 | -85.1154
0.84798 | -85.9366
0.8517 | -86.2293
0.85542 | -86.0142
0.85913 | -85.7956
0.86285 | -85.5192
0.86657 | -85.128
0.87029 | -84.7087
0.87401 | -84.5661
0.87773 | -84.9881
0.88145 | -85.5969
0.88517 | -86.0024
0.88889 | -86.2142
0.89261 | -86.1825
0.89633 | -86.1439
0.90005 | -86.4111
0.90377 | -86.6182
0.90748 | -86.4043
0.9112 | -86.274
0.91492 | -86.1761
0.91864 | -85.7115
0.92236 | -85.7209
0.92608 | -86.5211
0.9298 | -86.411
0.93352 | -85.5796
0.93724 | -85.3068
0.94096 | -85.4401
0.94468 | -85.5666
0.9484 | -85.8144
0.95212 | -86.2752
0.95583 | -86.274
0.95955 | -86.483
0.96327 | -86.6949
0.96699 | -86.3383
0.97071 | -86.0235
0.97443 | -86.3034
0.97815 | -87.1439
0.98187 | -87.7755
0.98559 | -87.5669
0.98931 | -86.8572
0.99303 | -86.6446
0.99675 | -86.7484
1.0005 | -86.753
1.0042 | -86.8904
1.0079 | -87.0895
1.0116 | -87.1744
1.0153 | -86.8749
1.0191 | -86.8831
1.0228 | -87.1832
1.0265 | -87.4715
1.0302 | -87.1908
1.0339 | -86.8041
1.0377 | -86.9155
1.0414 | -87.2084
1.0451 | -87.1964
1.0488 | -87.1407
1.0525 | -87.43
1.0563 | -87.2615
1.06 | -87.0725
1.0637 | -87.3953
1.0674 | -87.6491
1.0711 | -87.16
1.0748 | -86.6783
1.0786 | -86.3358
1.0823 | -86.2953
1.086 | -86.5031
1.0897 | -86.7833
1.0934 | -86.9777
1.0972 | -87.1895
1.1009 | -87.5211
1.1046 | -87.8208
1.1083 | -88.0053
1.112 | -88.0987
1.1158 | -88.1039
1.1195 | -88.1317
1.1232 | -88.1668
1.1269 | -87.9565
1.1306 | -87.5794
1.1344 | -87.4364
1.1381 | -87.7124
1.1418 | -88.2578
1.1455 | -88.7427
1.1492 | -88.9157
1.153 | -88.7356
1.1567 | -88.456
1.1604 | -88.3432
1.1641 | -88.3313
1.1678 | -88.2284
1.1715 | -87.9141
1.1753 | -87.3856
1.179 | -86.9586
1.1827 | -86.9071
1.1864 | -87.1642
1.1901 | -87.5244
1.1939 | -87.7666
1.1976 | -87.7873
1.2013 | -87.7718
1.205 | -87.967
1.2087 | -88.3908
1.2125 | -88.663
1.2162 | -88.4875
1.2199 | -88.2652
1.2236 | -88.3644
1.2273 | -88.6517
1.2311 | -88.8303
1.2348 | -88.8728
1.2385 | -88.9543
1.2422 | -89.1834
1.2459 | -89.3873
1.2497 | -89.2182
1.2534 | -88.7676
1.2571 | -88.4447
1.2608 | -88.3833
1.2645 | -88.4811
1.2682 | -88.5946
1.272 | -88.7136
1.2757 | -88.8816
1.2794 | -89.0112
1.2831 | -89.0139
1.2868 | -88.9506
1.2906 | -88.8981
1.2943 | -88.9432
1.298 | -89.1672
1.3017 | -89.5647
1.3054 | -89.9703
1.3092 | -90.1704
1.3129 | -90.0455
1.3166 | -89.7189
1.3203 | -89.5086
1.324 | -89.4567
1.3278 | -89.3629
1.3315 | -89.3814
1.3352 | -89.715
1.3389 | -90.0816
1.3426 | -90.1154
1.3464 | -89.8978
1.3501 | -89.6896
1.3538 | -89.5422
1.3575 | -89.4517
1.3612 | -89.5104
1.3649 | -89.6792
1.3687 | -89.8365
1.3724 | -89.9725
1.3761 | -90.0429
1.3798 | -89.9755
1.3835 | -89.7892
1.3873 | -89.6249
1.391 | -89.6092
1.3947 | -89.6591
1.3984 | -89.6068
1.4021 | -89.5268
1.4059 | -89.678
1.4096 | -90.0882
1.4133 | -90.3432
1.417 | -90.1008
1.4207 | -89.7084
1.4245 | -89.5944
1.4282 | -89.8143
1.4319 | -90.218
1.4356 | -90.5423
1.4393 | -90.6294
1.443 | -90.5982
1.4468 | -90.6152
1.4505 | -90.6665
1.4542 | -90.6701
1.4579 | -90.6057
1.4616 | -90.4537
1.4654 | -90.3647
1.4691 | -90.5011
1.4728 | -90.6803
1.4765 | -90.7359
1.4802 | -90.7894
1.484 | -90.9821
1.4877 | -91.3246
1.4914 | -91.6172
1.4951 | -91.6108
1.4988 | -91.4094
1.5026 | -91.3063
1.5063 | -91.3263
1.51 | -91.2643
1.5137 | -91.1157
1.5174 | -91.041
1.5212 | -91.0564
1.5249 | -91.0827
1.5286 | -91.0439
1.5323 | -90.8714
1.536 | -90.6428
1.5397 | -90.6045
1.5435 | -90.9118
1.5472 | -91.28
1.5509 | -91.2746
1.5546 | -90.9652
1.5583 | -90.659
1.5621 | -90.4847
1.5658 | -90.3971
1.5695 | -90.2808
1.5732 | -90.1577
1.5769 | -90.123
1.5807 | -90.1439
1.5844 | -90.214
1.5881 | -90.4796
1.5918 | -90.9026
1.5955 | -91.1087
1.5993 | -90.9714
1.603 | -90.9064
1.6067 | -91.1394
1.6104 | -91.2704
1.6141 | -91.3659
1.6179 | -91.4854
1.6216 | -91.6891
1.6253 | -91.9149
1.629 | -92.0498
1.6327 | -92.0894
1.6364 | -92.156
1.6402 | -92.3184
1.6439 | -92.4996
1.6476 | -92.5307
1.6513 | -92.3174
1.655 | -91.953
1.6588 | -91.6041
1.6625 | -91.3841
1.6662 | -91.2915
1.6699 | -91.2635
1.6736 | -91.2606
1.6774 | -91.283
1.6811 | -91.3548
1.6848 | -91.479
1.6885 | -91.6056
1.6922 | -91.7063
1.696 | -91.8047
1.6997 | -91.9437
1.7034 | -92.1138
1.7071 | -92.2149
1.7108 | -92.1893
1.7146 | -92.09
1.7183 | -92.0405
1.722 | -92.0756
1.7257 | -92.0702
1.7294 | -91.919
1.7331 | -91.6847
1.7369 | -91.483
1.7406 | -91.3934
1.7443 | -91.4459
1.748 | -91.6147
1.7517 | -91.781
1.7555 | -91.8756
1.7592 | -91.9561
1.7629 | -92.0606
1.7666 | -92.1106
1.7703 | -92.0548
1.7741 | -91.9495
1.7778 | -91.8663
1.7815 | -91.83
1.7852 | -91.7934
1.7889 | -91.7643
1.7927 | -91.832
1.7964 | -92.0424
1.8001 | -92.3303
1.8038 | -92.5984
1.8075 | -92.8077
1.8113 | -92.9166
1.815 | -92.8623
1.8187 | -92.7297
1.8224 | -92.6592
1.8261 | -92.7234
1.8298 | -92.881
1.8336 | -93.0292
1.8373 | -93.0922
1.841 | -93.058
1.8447 | -92.9728
1.8484 | -92.924
1.8522 | -92.9283
1.8559 | -92.9432
1.8596 | -92.9391
1.8633 | -92.939
1.867 | -93.0091
1.8708 | -93.1778
1.8745 | -93.3289
1.8782 | -93.327
1.8819 | -93.1411
1.8856 | -92.8955
1.8894 | -92.7025
1.8931 | -92.58
1.8968 | -92.5088
1.9005 | -92.4723
1.9042 | -92.4866
1.908 | -92.5706
1.9117 | -92.6921
1.9154 | -92.7893
1.9191 | -92.8079
1.9228 | -92.7729
1.9265 | -92.7559
1.9303 | -92.7471
1.934 | -92.7068
1.9377 | -92.6454
1.9414 | -92.6349
1.9451 | -92.6891
1.9489 | -92.7665
1.9526 | -92.8437
1.9563 | -92.9262
1.96 | -92.9756
1.9637 | -92.9491
1.9675 | -92.8305
1.9712 | -92.6978
1.9749 | -92.6194
1.9786 | -92.6026
1.9823 | -92.6036
1.9861 | -92.6221
1.9898 | -92.6975
1.9935 | -92.825
1.9972 | -92.9322
2.0009 | -92.9779
2.0046 | -92.931
2.0084 | -92.8178
2.0121 | -92.6965
2.0158 | -92.623
2.0195 | -92.6274
2.0232 | -92.7445
2.027 | -92.9896
2.0307 | -93.3093
2.0344 | -93.5435
2.0381 | -93.5346
2.0418 | -93.2955
2.0456 | -93.0528
2.0493 | -92.9864
2.053 | -93.0905
2.0567 | -93.2898
2.0604 | -93.5497
2.0642 | -93.8651
2.0679 | -94.1359
2.0716 | -94.173
2.0753 | -93.9175
2.079 | -93.4888
2.0828 | -93.0279
2.0865 | -92.6262
2.0902 | -92.3391
2.0939 | -92.2142
2.0976 | -92.2627
2.1013 | -92.4702
2.1051 | -92.7732
2.1088 | -93.0833
2.1125 | -93.3227
2.1162 | -93.4662
2.1199 | -93.5539
2.1237 | -93.6383
2.1274 | -93.7398
2.1311 | -93.8153
2.1348 | -93.7718
2.1385 | -93.5945
2.1423 | -93.3545
2.146 | -93.2391
2.1497 | -93.1863
2.1534 | -93.2419
2.1571 | -93.3791
2.1609 | -93.5323
2.1646 | -93.6315
2.1683 | -93.673
2.172 | -93.6926
2.1757 | -93.7174
2.1795 | -93.7523
2.1832 | -93.7983
2.1869 | -93.8503
2.1906 | -93.8675
2.1943 | -93.8115
2.198 | -93.696
2.2018 | -93.5709
2.2055 | -93.4919
2.2092 | -93.4746
2.2129 | -93.5114
2.2166 | -93.5471
2.2204 | -93.5493
2.2241 | -93.5308
2.2278 | -93.5082
2.2315 | -93.5052
2.2352 | -93.5262
2.239 | -93.5641
2.2427 | -93.6305
2.2464 | -93.7104
2.2501 | -93.7713
2.2538 | -93.7576
2.2576 | -93.6896
2.2613 | -93.6203
2.265 | -93.5813
2.2687 | -93.5869
2.2724 | -93.6228
2.2762 | -93.6833
2.2799 | -93.7935
2.2836 | -93.9525
2.2873 | -94.1307
2.291 | -94.2949
2.2947 | -94.4308
2.2985 | -94.4873
2.3022 | -94.4222
2.3059 | -94.2506
2.3096 | -94.032
2.3133 | -93.8218
2.3171 | -93.6728
2.3208 | -93.5889
2.3245 | -93.5375
2.3282 | -93.4927
2.3319 | -93.4214
2.3357 | -93.3271
2.3394 | -93.2395
2.3431 | -93.1936
2.3468 | -93.2151
2.3505 | -93.3323
2.3543 | -93.5731
2.358 | -93.8815
2.3617 | -94.1609
2.3654 | -94.3341
2.3691 | -94.3633
2.3729 | -94.313
2.3766 | -94.2732
2.3803 | -94.3098
2.384 | -94.4338
2.3877 | -94.6108
2.3914 | -94.7696
2.3952 | -94.8559
2.3989 | -94.8594
2.4026 | -94.7943
2.4063 | -94.6729
2.41 | -94.52
2.4138 | -94.3507
2.4175 | -94.203
2.4212 | -94.1379
2.4249 | -94.208
2.4286 | -94.4029
2.4324 | -94.6364
2.4361 | -94.8074
2.4398 | -94.8576
2.4435 | -94.7761
2.4472 | -94.5806
2.451 | -94.3206
2.4547 | -94.0694
2.4584 | -93.8955
2.4621 | -93.8422
2.4658 | -93.9212
2.4695 | -94.1189
2.4733 | -94.3966
2.477 | -94.6759
2.4807 | -94.8755
2.4844 | -94.9412
2.4881 | -94.8896
2.4919 | -94.7639
2.4956 | -94.6324
2.4993 | -94.5369
2.503 | -94.52
2.5067 | -94.5925
2.5105 | -94.7343
2.5142 | -94.9067
2.5179 | -95.0628
2.5216 | -95.1985
2.5253 | -95.3252
2.5291 | -95.4164
2.5328 | -95.4696
2.5365 | -95.4829
2.5402 | -95.4936
2.5439 | -95.5044
2.5477 | -95.4985
2.5514 | -95.4433
2.5551 | -95.3088
2.5588 | -95.1287
2.5625 | -94.9787
2.5662 | -94.9102
2.57 | -94.9294
2.5737 | -94.9723
2.5774 | -95.0009
2.5811 | -94.9585
2.5848 | -94.8268
2.5886 | -94.6002
2.5923 | -94.3192
2.596 | -94.0929
2.5997 | -93.9967
2.6034 | -94.0564
2.6072 | -94.2502
2.6109 | -94.5211
2.6146 | -94.8334
2.6183 | -95.1497
2.622 | -95.4314
2.6258 | -95.612
2.6295 | -95.655
2.6332 | -95.5788
2.6369 | -95.4152
2.6406 | -95.221
2.6444 | -95.0299
2.6481 | -94.8962
2.6518 | -94.8707
2.6555 | -94.9798
2.6592 | -95.1778
2.6629 | -95.3668
2.6667 | -95.4503
2.6704 | -95.4108
2.6741 | -95.3199
2.6778 | -95.2754
2.6815 | -95.4069
2.6853 | -95.5032
2.689 | -95.6101
2.6927 | -95.6959
2.6964 | -95.728
2.7001 | -95.6797
2.7039 | -95.5798
2.7076 | -95.4623
2.7113 | -95.3592
2.715 | -95.3021
2.7187 | -95.3054
2.7225 | -95.3537
2.7262 | -95.407
2.7299 | -95.4459
2.7336 | -95.4462
2.7373 | -95.3904
2.7411 | -95.275
2.7448 | -95.138
2.7485 | -95.0299
2.7522 | -94.982
2.7559 | -95.0153
2.7596 | -95.1313
2.7634 | -95.3282
2.7671 | -95.5771
2.7708 | -95.8238
2.7745 | -96.0089
2.7782 | -96.0773
2.782 | -96.0484
2.7857 | -95.9795
2.7894 | -95.9214
2.7931 | -95.8801
2.7968 | -95.843
2.8006 | -95.8123
2.8043 | -95.7805
2.808 | -95.759
2.8117 | -95.7552
2.8154 | -95.7739
2.8192 | -95.8388
2.8229 | -95.9235
2.8266 | -95.9981
2.8303 | -96.0286
2.834 | -95.9848
2.8378 | -95.8826
2.8415 | -95.7627
2.8452 | -95.6648
2.8489 | -95.6133
2.8526 | -95.6228
2.8563 | -95.6832
2.8601 | -95.7722
2.8638 | -95.8731
2.8675 | -95.9532
2.8712 | -96.0114
2.8749 | -96.054
2.8787 | -96.0989
2.8824 | -96.1551
2.8861 | -96.1886
2.8898 | -96.1727
2.8935 | -96.0975
2.8973 | -95.9504
2.901 | -95.7737
2.9047 | -95.619
2.9084 | -95.526
2.9121 | -95.511
2.9159 | -95.5634
2.9196 | -95.6497
2.9233 | -95.7311
2.927 | -95.7813
2.9307 | -95.8032
2.9344 | -95.8173
2.9382 | -95.8523
2.9419 | -95.9299
2.9456 | -96.0538
2.9493 | -96.1618
2.953 | -96.2045
2.9568 | -96.1678
2.9605 | -96.0653
2.9642 | -95.9444
2.9679 | -95.8489
2.9716 | -95.8128
2.9754 | -95.8399
2.9791 | -95.8997
2.9828 | -95.9611
2.9865 | -95.9841
2.9902 | -95.9467
2.994 | -95.8661
2.9977 | -95.7719
3.0014 | -95.6846
3.0051 | -95.6264
3.0088 | -95.6213
3.0126 | -95.6784
3.0163 | -95.794
3.02 | -95.9419
3.0237 | -96.1069
3.0274 | -96.2673
3.0311 | -96.4059
3.0349 | -96.5186
3.0386 | -96.582
3.0423 | -96.6035
3.046 | -96.5739
3.0497 | -96.4978
3.0535 | -96.3862
3.0572 | -96.2592
3.0609 | -96.1524
3.0646 | -96.097
3.0683 | -96.1183
3.0721 | -96.2046
3.0758 | -96.3397
3.0795 | -96.4694
3.0832 | -96.5579
3.0869 | -96.5806
3.0907 | -96.5335
3.0944 | -96.4564
3.0981 | -96.3687
3.1018 | -96.2939
3.1055 | -96.2359
3.1093 | -96.2161
3.113 | -96.2518
3.1167 | -96.3427
3.1204 | -96.4762
3.1241 | -96.6079
3.1278 | -96.7108
3.1316 | -96.7684
3.1353 | -96.7971
3.139 | -96.7895
3.1427 | -96.7573
3.1464 | -96.7133
3.1502 | -96.6683
3.1539 | -96.6576
3.1576 | -96.679
3.1613 | -96.7269
3.165 | -96.7665
3.1688 | -96.7619
3.1725 | -96.7096
3.1762 | -96.6215
3.1799 | -96.5377
3.1836 | -96.5136
3.1874 | -96.5845
3.1911 | -96.7459
3.1948 | -96.9598
3.1985 | -97.1554
3.2022 | -97.2653
3.206 | -97.2724
3.2097 | -97.2233
3.2134 | -97.1758
3.2171 | -97.2253
3.2208 | -97.2603
3.2245 | -97.302
3.2283 | -97.3428
3.232 | -97.3688
3.2357 | -97.3717
3.2394 | -97.3632
3.2431 | -97.362
3.2469 | -97.3647
3.2506 | -97.3637
3.2543 | -97.3499
3.258 | -97.3125
3.2617 | -97.2408
3.2655 | -97.1653
3.2692 | -97.1001
3.2729 | -97.048
3.2766 | -97.0275
3.2803 | -97.0341
3.2841 | -97.0587
3.2878 | -97.0835
3.2915 | -97.1035
3.2952 | -97.1203
3.2989 | -97.1328
3.3026 | -97.1734
3.3064 | -97.2538
3.3101 | -97.3587
3.3138 | -97.4671
3.3175 | -97.5467
3.3212 | -97.5622
3.325 | -97.5138
3.3287 | -97.4246
3.3324 | -97.3104
3.3361 | -97.2111
3.3398 | -97.1433
3.3436 | -97.1197
3.3473 | -97.1443
3.351 | -97.2141
3.3547 | -97.3303
3.3584 | -97.4692
3.3622 | -97.6062
3.3659 | -97.7139
3.3696 | -97.7652
3.3733 | -97.7551
3.377 | -97.6927
3.3808 | -97.6032
3.3845 | -97.512
3.3882 | -97.433
3.3919 | -97.3668
3.3956 | -97.3239
3.3993 | -97.306
3.4031 | -97.3034
3.4068 | -97.32
3.4105 | -97.3429
3.4142 | -97.3707
3.4179 | -97.3944
3.4217 | -97.4033
3.4254 | -97.3891
3.4291 | -97.3553
3.4328 | -97.307
3.4365 | -97.2508
3.4403 | -97.1975
3.444 | -97.1643
3.4477 | -97.151
3.4514 | -97.1477
3.4551 | -97.162
3.4589 | -97.2048
3.4626 | -97.2769
3.4663 | -97.3853
3.47 | -97.5206
3.4737 | -97.6651
3.4775 | -97.796
3.4812 | -97.8896
3.4849 | -97.9447
3.4886 | -97.968
3.4923 | -97.9666
3.496 | -97.9453
3.4998 | -97.9314
3.5035 | -97.9292
3.5072 | -97.9466
3.5109 | -97.9718
3.5146 | -97.989
3.5184 | -97.9976
3.5221 | -97.9966
3.5258 | -97.9744
3.5295 | -97.9241
3.5332 | -97.843
3.537 | -97.7445
3.5407 | -97.6378
3.5444 | -97.5484
3.5481 | -97.4978
3.5518 | -97.4792
3.5556 | -97.4575
3.5593 | -97.432
3.563 | -97.3772
3.5667 | -97.317
3.5704 | -97.2879
3.5742 | -97.298
3.5779 | -97.3549
3.5816 | -97.4596
3.5853 | -97.5958
3.589 | -97.7262
3.5927 | -97.8479
3.5965 | -97.951
3.6002 | -98.0169
3.6039 | -98.0377
3.6076 | -98.0286
3.6113 | -98.0083
3.6151 | -97.9855
3.6188 | -97.9807
3.6225 | -97.9801
3.6262 | -97.9871
3.6299 | -98.0013
3.6337 | -98.0079
3.6374 | -97.9997
3.6411 | -97.979
3.6448 | -97.9521
3.6485 | -97.9214
3.6523 | -97.9069
3.656 | -97.9108
3.6597 | -97.9221
3.6634 | -97.9348
3.6671 | -97.9352
3.6709 | -97.9297
3.6746 | -97.9168
3.6783 | -97.9017
3.682 | -97.8798
3.6857 | -97.8617
3.6894 | -97.8429
3.6932 | -97.8201
3.6969 | -97.8006
3.7006 | -97.7985
3.7043 | -97.8032
3.708 | -97.8101
3.7118 | -97.826
3.7155 | -97.8478
3.7192 | -97.8726
3.7229 | -97.9127
3.7266 | -97.9584
3.7304 | -98.0127
3.7341 | -98.0729
3.7378 | -98.1341
3.7415 | -98.1929
3.7452 | -98.2545
3.749 | -98.3208
3.7527 | -98.3535
3.7564 | -98.389
3.7601 | -98.3899
3.7638 | -98.3621
3.7675 | -98.3143
3.7713 | -98.2519
3.775 | -98.191
3.7787 | -98.1494
3.7824 | -98.1304
3.7861 | -98.1467
3.7899 | -98.1926
3.7936 | -98.2451
3.7973 | -98.3024
3.801 | -98.3447
3.8047 | -98.3565
3.8085 | -98.338
3.8122 | -98.296
3.8159 | -98.244
3.8196 | -98.1974
3.8233 | -98.1784
3.8271 | -98.1791
3.8308 | -98.2102
3.8345 | -98.2618
3.8382 | -98.319
3.8419 | -98.3755
3.8457 | -98.4226
3.8494 | -98.4482
3.8531 | -98.4581
3.8568 | -98.4711
3.8605 | -98.4694
3.8642 | -98.465
3.868 | -98.4668
3.8717 | -98.4625
3.8754 | -98.4465
3.8791 | -98.4111
3.8828 | -98.3653
3.8866 | -98.3158
3.8903 | -98.2837
3.894 | -98.2614
3.8977 | -98.261
3.9014 | -98.291
3.9052 | -98.3294
3.9089 | -98.3791
3.9126 | -98.4217
3.9163 | -98.439
3.92 | -98.4269
3.9238 | -98.3978
3.9275 | -98.3484
3.9312 | -98.2901
3.9349 | -98.2606
3.9386 | -98.2755
3.9424 | -98.3364
3.9461 | -98.4365
3.9498 | -98.5663
3.9535 | -98.7049
3.9572 | -98.8427
3.9609 | -98.9748
3.9647 | -99.0722
3.9684 | -99.133
3.9721 | -99.1464
3.9758 | -99.1179
3.9795 | -99.0578
3.9833 | -98.9747
3.987 | -98.8846
3.9907 | -98.8218
3.9944 | -98.7957
3.9981 | -98.7945
4.0019 | -98.8241
4.0056 | -98.8696
4.0093 | -98.92
4.013 | -98.9619
4.0167 | -98.9957
4.0205 | -99.018
4.0242 | -99.0323
4.0279 | -99.0432
4.0316 | -99.0501
4.0353 | -99.0426
4.0391 | -99.022
4.0428 | -98.9881
4.0465 | -98.9428
4.0502 | -98.8883
4.0539 | -98.83
4.0576 | -98.7775
4.0614 | -98.7443
4.0651 | -98.7322
4.0688 | -98.7335
4.0725 | -98.7318
4.0762 | -98.7193
4.08 | -98.6795
4.0837 | -98.6263
4.0874 | -98.5521
4.0911 | -98.4724
4.0948 | -98.382
4.0986 | -98.2945
4.1023 | -98.2283
4.106 | -98.1842
4.1097 | -98.168
4.1134 | -98.1886
4.1172 | -98.2383
4.1209 | -98.3068
4.1246 | -98.385
4.1283 | -98.4607
4.132 | -98.5235
4.1358 | -98.5646
4.1395 | -98.5778
4.1432 | -98.5608
4.1469 | -98.5102
4.1506 | -98.4369
4.1543 | -98.3596
4.1581 | -98.3015
4.1618 | -98.275
4.1655 | -98.2876
4.1692 | -98.3383
4.1729 | -98.4147
4.1767 | -98.4929
4.1804 | -98.5697
4.1841 | -98.6436
4.1878 | -98.7076
4.1915 | -98.7632
4.1953 | -98.8162
4.199 | -98.8611
4.2027 | -98.8927
4.2064 | -98.9155
4.2101 | -98.9181
4.2139 | -98.9163
4.2176 | -98.9056
4.2213 | -98.8999
4.225 | -98.9019
4.2287 | -98.9137
4.2325 | -98.9435
4.2362 | -98.9825
4.2399 | -99.0187
4.2436 | -99.0421
4.2473 | -99.0578
4.251 | -99.0731
4.2548 | -99.0837
4.2585 | -99.0962
4.2622 | -99.1239
4.2659 | -99.1604
4.2696 | -99.2086
4.2734 | -99.2675
4.2771 | -99.3294
4.2808 | -99.3976
4.2845 | -99.4505
4.2882 | -99.4922
4.292 | -99.4875
4.2957 | -99.4961
4.2994 | -99.49
4.3031 | -99.4749
4.3068 | -99.4583
4.3106 | -99.4326
4.3143 | -99.4021
4.318 | -99.3759
4.3217 | -99.3511
4.3254 | -99.3189
4.3291 | -99.2643
4.3329 | -99.208
4.3366 | -99.1497
4.3403 | -99.0827
4.344 | -99.0143
4.3477 | -98.9505
4.3515 | -98.8976
4.3552 | -98.8551
4.3589 | -98.8303
4.3626 | -98.8181
4.3663 | -98.8115
4.3701 | -98.807
4.3738 | -98.8156
4.3775 | -98.8251
4.3812 | -98.8384
4.3849 | -98.8601
4.3887 | -98.8849
4.3924 | -98.9238
4.3961 | -98.9812
4.3998 | -99.0538
4.4035 | -99.13
4.4073 | -99.2106
4.411 | -99.2992
4.4147 | -99.3722
4.4184 | -99.4188
4.4221 | -99.4302
4.4258 | -99.3968
4.4296 | -99.3207
4.4333 | -99.2171
4.437 | -99.1122
4.4407 | -99.0085
4.4444 | -98.916
4.4482 | -98.8548
4.4519 | -98.825
4.4556 | -98.8222
4.4593 | -98.8583
4.463 | -98.9234
4.4668 | -99.0034
4.4705 | -99.1089
4.4742 | -99.2224
4.4779 | -99.3318
4.4816 | -99.4279
4.4854 | -99.5095
4.4891 | -99.5667
4.4928 | -99.6014
4.4965 | -99.6189
4.5002 | -99.6212
4.504 | -99.5968
4.5077 | -99.5542
4.5114 | -99.5022
4.5151 | -99.4316
4.5188 | -99.3631
4.5225 | -99.3074
4.5263 | -99.2678
4.53 | -99.2431
4.5337 | -99.2489
4.5374 | -99.2848
4.5411 | -99.3312
4.5449 | -99.3924
4.5486 | -99.4529
4.5523 | -99.5066
4.556 | -99.5411
4.5597 | -99.5645
4.5635 | -99.5804
4.5672 | -99.5926
4.5709 | -99.6052
4.5746 | -99.6376
4.5783 | -99.6829
4.5821 | -99.7368
4.5858 | -99.7997
4.5895 | -99.8549
4.5932 | -99.9018
4.5969 | -99.9437
4.6007 | -99.9786
4.6044 | -99.9891
4.6081 | -99.9816
4.6118 | -99.9642
4.6155 | -99.9364
4.6192 | -99.9086
4.623 | -99.8914
4.6267 | -99.8969
4.6304 | -99.9165
4.6341 | -99.947
4.6378 | -99.9788
4.6416 | -100.0142
4.6453 | -100.0339
4.649 | -100.0412
4.6527 | -100.0419
4.6564 | -100.0469
4.6602 | -100.0631
4.6639 | -100.0892
4.6676 | -100.1242
4.6713 | -100.1591
4.675 | -100.1995
4.6788 | -100.2197
4.6825 | -100.215
4.6862 | -100.1956
4.6899 | -100.1455
4.6936 | -100.0707
4.6974 | -99.9824
4.7011 | -99.8946
4.7048 | -99.8193
4.7085 | -99.7707
4.7122 | -99.7508
4.7159 | -99.7718
4.7197 | -99.82
4.7234 | -99.8988
4.7271 | -99.9926
4.7308 | -100.0931
4.7345 | -100.1926
4.7383 | -100.2888
4.742 | -100.3774
4.7457 | -100.4402
4.7494 | -100.4988
4.7531 | -100.5398
4.7569 | -100.5592
4.7606 | -100.5709
4.7643 | -100.5659
4.768 | -100.5489
4.7717 | -100.5131
4.7755 | -100.4572
4.7792 | -100.3873
4.7829 | -100.3057
4.7866 | -100.2211
4.7903 | -100.1382
4.794 | -100.0622
4.7978 | -100.0019
4.8015 | -99.9682
4.8052 | -99.9673
4.8089 | -99.9967
4.8126 | -100.0474
4.8164 | -100.1142
4.8201 | -100.1648
4.8238 | -100.1939
4.8275 | -100.1544
4.8312 | -100.1386
4.835 | -100.1113
4.8387 | -100.0746
4.8424 | -100.0364
4.8461 | -100.0127
4.8498 | -100.0095
4.8536 | -100.0187
4.8573 | -100.0439
4.861 | -100.0851
4.8647 | -100.1389
4.8684 | -100.1929
4.8722 | -100.245
4.8759 | -100.2902
4.8796 | -100.3225
4.8833 | -100.3475
4.887 | -100.3701
4.8907 | -100.3938
4.8945 | -100.4249
4.8982 | -100.45
4.9019 | -100.4583
4.9056 | -100.4555
4.9093 | -100.4521
4.9131 | -100.4346
4.9168 | -100.4022
4.9205 | -100.3607
4.9242 | -100.321
4.9279 | -100.2857
4.9317 | -100.2621
4.9354 | -100.2456
4.9391 | -100.2395
4.9428 | -100.2445
4.9465 | -100.2557
4.9503 | -100.2714
4.954 | -100.296
4.9577 | -100.323
4.9614 | -100.3467
4.9651 | -100.3745
4.9689 | -100.4032
4.9726 | -100.4302
4.9763 | -100.4641
4.98 | -100.4977
4.9837 | -100.5214
4.9874 | -100.5351
4.9912 | -100.5488
4.9949 | -100.5646
4.9986 | -100.5672
5.0023 | -100.5593
5.006 | -100.5458
5.0098 | -100.528
5.0135 | -100.5151
5.0172 | -100.4983
5.0209 | -100.4863
5.0246 | -100.4829
5.0284 | -100.4863
5.0321 | -100.4895
5.0358 | -100.5003
5.0395 | -100.513
5.0432 | -100.5167
5.047 | -100.5088
5.0507 | -100.4873
5.0544 | -100.4579
5.0581 | -100.429
5.0618 | -100.3932
5.0656 | -100.3535
5.0693 | -100.3168
5.073 | -100.2853
5.0767 | -100.2711
5.0804 | -100.265
5.0841 | -100.2645
5.0879 | -100.2746
5.0916 | -100.2869
5.0953 | -100.302
5.099 | -100.3207
5.1027 | -100.341
5.1065 | -100.3559
5.1102 | -100.3646
5.1139 | -100.373
5.1176 | -100.3774
5.1213 | -100.3807
5.1251 | -100.3859
5.1288 | -100.3914
5.1325 | -100.3945
5.1362 | -100.3974
5.1399 | -100.412
5.1437 | -100.4293
5.1474 | -100.4505
5.1511 | -100.4755
5.1548 | -100.4997
5.1585 | -100.5263
5.1623 | -100.5511
5.166 | -100.5668
5.1697 | -100.5723
5.1734 | -100.5697
5.1771 | -100.559
5.1808 | -100.5407
5.1846 | -100.5209
5.1883 | -100.5036
5.192 | -100.4841
5.1957 | -100.4705
5.1994 | -100.4698
5.2032 | -100.4761
5.2069 | -100.4877
5.2106 | -100.5017
5.2143 | -100.5069
5.218 | -100.4974
5.2218 | -100.4729
5.2255 | -100.4274
5.2292 | -100.3665
5.2329 | -100.3062
5.2366 | -100.2525
5.2404 | -100.2203
5.2441 | -100.2135
5.2478 | -100.2324
5.2515 | -100.2737
5.2552 | -100.335
5.2589 | -100.4012
5.2627 | -100.4648
5.2664 | -100.5238
5.2701 | -100.5726
5.2738 | -100.6086
5.2775 | -100.6315
5.2813 | -100.6386
5.285 | -100.6321
5.2887 | -100.6056
5.2924 | -100.571
5.2961 | -100.5264
5.2999 | -100.4813
5.3036 | -100.4356
5.3073 | -100.4026
5.311 | -100.3876
5.3147 | -100.3959
5.3185 | -100.4263
5.3222 | -100.4808
5.3259 | -100.5503
5.3296 | -100.6221
5.3333 | -100.6937
5.3371 | -100.7618
5.3408 | -100.815
5.3445 | -100.8591
5.3482 | -100.9012
5.3519 | -100.9388
5.3556 | -100.969
5.3594 | -101.0027
5.3631 | -101.039
5.3668 | -101.0751
5.3705 | -101.1065
5.3742 | -101.1271
5.378 | -101.1457
5.3817 | -101.1523
5.3854 | -101.1537
5.3891 | -101.1458
5.3928 | -101.1245
5.3966 | -101.0953
5.4003 | -101.059
5.404 | -101.0064
5.4077 | -100.9551
5.4114 | -100.9094
5.4152 | -100.8632
5.4189 | -100.8148
5.4226 | -100.773
5.4263 | -100.7278
5.43 | -100.6887
5.4338 | -100.6573
5.4375 | -100.6323
5.4412 | -100.6099
5.4449 | -100.5781
5.4486 | -100.5497
5.4523 | -100.5239
5.4561 | -100.5038
5.4598 | -100.4917
5.4635 | -100.4817
5.4672 | -100.4814
5.4709 | -100.4903
5.4747 | -100.501
5.4784 | -100.5133
5.4821 | -100.5272
5.4858 | -100.548
5.4895 | -100.5798
5.4933 | -100.6122
5.497 | -100.6414
5.5007 | -100.672
5.5044 | -100.7111
5.5081 | -100.749
5.5119 | -100.7859
5.5156 | -100.8191
5.5193 | -100.8537
5.523 | -100.886
5.5267 | -100.9085
5.5305 | -100.9368
5.5342 | -100.9576
5.5379 | -100.9734
5.5416 | -100.9912
5.5453 | -101.0064
5.549 | -101.0287
5.5528 | -101.057
5.5565 | -101.0818
5.5602 | -101.1059
5.5639 | -101.13
5.5676 | -101.1488
5.5714 | -101.151
5.5751 | -101.1389
5.5788 | -101.1281
5.5825 | -101.1178
5.5862 | -101.1081
5.59 | -101.0989
5.5937 | -101.0942
5.5974 | -101.0933
5.6011 | -101.1047
5.6048 | -101.1235
5.6086 | -101.1492
5.6123 | -101.1781
5.616 | -101.2029
5.6197 | -101.229
5.6234 | -101.2463
5.6272 | -101.2512
5.6309 | -101.2474
5.6346 | -101.2333
5.6383 | -101.199
5.642 | -101.1612
5.6457 | -101.1176
5.6495 | -101.0664
5.6532 | -101.0217
5.6569 | -100.9868
5.6606 | -100.9667
5.6643 | -100.9568
5.6681 | -100.9657
5.6718 | -100.9873
5.6755 | -101.0225
5.6792 | -101.0689
5.6829 | -101.1242
5.6867 | -101.1733
5.6904 | -101.2251
5.6941 | -101.2836
5.6978 | -101.3422
5.7015 | -101.3938
5.7053 | -101.4345
5.709 | -101.4707
5.7127 | -101.5
5.7164 | -101.5217
5.7201 | -101.5247
5.7238 | -101.5231
5.7276 | -101.5092
5.7313 | -101.492
5.735 | -101.4711
5.7387 | -101.4393
5.7424 | -101.4057
5.7462 | -101.3691
5.7499 | -101.3277
5.7536 | -101.2889
5.7573 | -101.2554
5.761 | -101.2261
5.7648 | -101.2009
5.7685 | -101.1773
5.7722 | -101.1673
5.7759 | -101.1625
5.7796 | -101.1604
5.7834 | -101.1597
5.7871 | -101.1628
5.7908 | -101.1745
5.7945 | -101.1907
5.7982 | -101.2046
5.802 | -101.2108
5.8057 | -101.2078
5.8094 | -101.1986
5.8131 | -101.1829
5.8168 | -101.1552
5.8205 | -101.1279
5.8243 | -101.104
5.828 | -101.0814
5.8317 | -101.0733
5.8354 | -101.0798
5.8391 | -101.0895
5.8429 | -101.1055
5.8466 | -101.1253
5.8503 | -101.1483
5.854 | -101.1717
5.8577 | -101.1773
5.8615 | -101.1739
5.8652 | -101.1649
5.8689 | -101.1556
5.8726 | -101.1412
5.8763 | -101.1209
5.8801 | -101.0988
5.8838 | -101.0839
5.8875 | -101.0786
5.8912 | -101.0812
5.8949 | -101.0833
5.8987 | -101.1108
5.9024 | -101.1266
5.9061 | -101.154
5.9098 | -101.1861
5.9135 | -101.226
5.9172 | -101.2788
5.921 | -101.3292
5.9247 | -101.3823
5.9284 | -101.4463
5.9321 | -101.5092
5.9358 | -101.5716
5.9396 | -101.6301
5.9433 | -101.6787
5.947 | -101.7233
5.9507 | -101.7495
5.9544 | -101.7579
5.9582 | -101.7538
5.9619 | -101.7357
5.9656 | -101.7034
5.9693 | -101.6656
5.973 | -101.6178
5.9768 | -101.5699
5.9805 | -101.5278
5.9842 | -101.4907
5.9879 | -101.4505
5.9916 | -101.4194
5.9954 | -101.3949
5.9991 | -101.368
6.0028 | -101.3407
6.0065 | -101.3164
6.0102 | -101.291
6.0139 | -101.2734
6.0177 | -101.2609
6.0214 | -101.2541
6.0251 | -101.2591
6.0288 | -101.2758
6.0325 | -101.3127
6.0363 | -101.3719
6.04 | -101.4514
6.0437 | -101.5533
6.0474 | -101.6662
6.0511 | -101.7801
6.0549 | -101.8936
6.0586 | -102.0052
6.0623 | -102.1045
6.066 | -102.1841
6.0697 | -102.2384
6.0735 | -102.2642
6.0772 | -102.2658
6.0809 | -102.2444
6.0846 | -102.1948
6.0883 | -102.1332
6.0921 | -102.0698
6.0958 | -102.0035
6.0995 | -101.9349
6.1032 | -101.8717
6.1069 | -101.8142
6.1106 | -101.7637
6.1144 | -101.7188
6.1181 | -101.6835
6.1218 | -101.6512
6.1255 | -101.6174
6.1292 | -101.5862
6.133 | -101.5556
6.1367 | -101.5284
6.1404 | -101.518
6.1441 | -101.5218
6.1478 | -101.5322
6.1516 | -101.5514
6.1553 | -101.5729
6.159 | -101.5986
6.1627 | -101.6258
6.1664 | -101.6508
6.1702 | -101.6685
6.1739 | -101.6794
6.1776 | -101.6828
6.1813 | -101.6713
6.185 | -101.6592
6.1887 | -101.648
6.1925 | -101.6358
6.1962 | -101.6355
6.1999 | -101.6475
6.2036 | -101.6674
6.2073 | -101.692
6.2111 | -101.7203
6.2148 | -101.7495
6.2185 | -101.7787
6.2222 | -101.809
6.2259 | -101.8243
6.2297 | -101.8235
6.2334 | -101.8116
6.2371 | -101.7868
6.2408 | -101.7558
6.2445 | -101.7254
6.2483 | -101.6973
6.252 | -101.6744
6.2557 | -101.6579
6.2594 | -101.652
6.2631 | -101.6584
6.2669 | -101.6801
6.2706 | -101.7152
6.2743 | -101.7675
6.278 | -101.8194
6.2817 | -101.8777
6.2854 | -101.9479
6.2892 | -102.0128
6.2929 | -102.0708
6.2966 | -102.1269
6.3003 | -102.1776
6.304 | -102.213
6.3078 | -102.2379
6.3115 | -102.2451
6.3152 | -102.2362
6.3189 | -102.2255
6.3226 | -102.2026
6.3264 | -102.1659
6.3301 | -102.1301
6.3338 | -102.0881
6.3375 | -102.0423
6.3412 | -102.0024
6.345 | -101.9642
6.3487 | -101.9281
6.3524 | -101.8903
6.3561 | -101.849
6.3598 | -101.8056
6.3636 | -101.761
6.3673 | -101.7162
6.371 | -101.6785
6.3747 | -101.6487
6.3784 | -101.619
6.3821 | -101.5991
6.3859 | -101.5952
6.3896 | -101.6118
6.3933 | -101.6442
6.397 | -101.6894
6.4007 | -101.7478
6.4045 | -101.8167
6.4082 | -101.8873
6.4119 | -101.9562
6.4156 | -102.0211
6.4193 | -102.0697
6.4231 | -102.103
6.4268 | -102.1291
6.4305 | -102.1338
6.4342 | -102.1258
6.4379 | -102.1132
6.4417 | -102.1095
6.4454 | -102.1017
6.4491 | -102.0983
6.4528 | -102.0993
6.4565 | -102.102
6.4603 | -102.1031
6.464 | -102.1018
6.4677 | -102.1027
6.4714 | -102.0978
6.4751 | -102.0938
6.4788 | -102.0868
6.4826 | -102.0721
6.4863 | -102.0482
6.49 | -102.0284
6.4937 | -102.0131
6.4974 | -102.0024
6.5012 | -101.9984
6.5049 | -102.0067
6.5086 | -102.0222
6.5123 | -102.0494
6.516 | -102.0812
6.5198 | -102.1094
6.5235 | -102.128
6.5272 | -102.1442
6.5309 | -102.151
6.5346 | -102.1492
6.5384 | -102.1406
6.5421 | -102.1166
6.5458 | -102.0879
6.5495 | -102.056
6.5532 | -102.0237
6.5569 | -101.9857
6.5607 | -101.9521
6.5644 | -101.9276
6.5681 | -101.9076
6.5718 | -101.8956
6.5755 | -101.8883
6.5793 | -101.8845
6.583 | -101.8955
6.5867 | -101.9168
6.5904 | -101.9409
6.5941 | -101.9659
6.5979 | -101.9923
6.6016 | -102.0175
6.6053 | -102.04
6.609 | -102.0621
6.6127 | -102.0749
6.6165 | -102.084
6.6202 | -102.0877
6.6239 | -102.0877
6.6276 | -102.0759
6.6313 | -102.0585
6.6351 | -102.0375
6.6388 | -102.0178
6.6425 | -102.0035
6.6462 | -101.9861
6.6499 | -101.9667
6.6536 | -101.9508
6.6574 | -101.9432
6.6611 | -101.939
6.6648 | -101.9378
6.6685 | -101.9453
6.6722 | -101.9537
6.676 | -101.9641
6.6797 | -101.9759
6.6834 | -101.9802
6.6871 | -101.9849
6.6908 | -101.9924
6.6946 | -101.9983
6.6983 | -101.9993
6.702 | -102.0008
6.7057 | -101.994
6.7094 | -101.9873
6.7132 | -101.99
6.7169 | -101.9955
6.7206 | -102.0016
6.7243 | -102.0137
6.728 | -102.0342
6.7318 | -102.0592
6.7355 | -102.08
6.7392 | -102.1074
6.7429 | -102.1345
6.7466 | -102.164
6.7503 | -102.1933
6.7541 | -102.2244
6.7578 | -102.2587
6.7615 | -102.2896
6.7652 | -102.3141
6.7689 | -102.3426
6.7727 | -102.3762
6.7764 | -102.4084
6.7801 | -102.4395
6.7838 | -102.4641
6.7875 | -102.4742
6.7913 | -102.4703
6.795 | -102.4611
6.7987 | -102.4358
6.8024 | -102.3979
6.8061 | -102.3596
6.8099 | -102.3104
6.8136 | -102.255
6.8173 | -102.1995
6.821 | -102.1501
6.8247 | -102.1035
6.8285 | -102.0652
6.8322 | -102.0402
6.8359 | -102.0361
6.8396 | -102.0436
6.8433 | -102.0736
6.847 | -102.1236
6.8508 | -102.1875
6.8545 | -102.2546
6.8582 | -102.3277
6.8619 | -102.4086
6.8656 | -102.4912
6.8694 | -102.5579
6.8731 | -102.6106
6.8768 | -102.6466
6.8805 | -102.6674
6.8842 | -102.6732
6.888 | -102.6639
6.8917 | -102.6459
6.8954 | -102.6137
6.8991 | -102.579
6.9028 | -102.5402
6.9066 | -102.4984
6.9103 | -102.455
6.914 | -102.4203
6.9177 | -102.3882
6.9214 | -102.3619
6.9252 | -102.3416
6.9289 | -102.3295
6.9326 | -102.3202
6.9363 | -102.3155
6.94 | -102.3195
6.9437 | -102.3229
6.9475 | -102.3243
6.9512 | -102.3234
6.9549 | -102.334
6.9586 | -102.3458
6.9623 | -102.3558
6.9661 | -102.3618
6.9698 | -102.3756
6.9735 | -102.3998
6.9772 | -102.4243
6.9809 | -102.4544
6.9847 | -102.481
6.9884 | -102.5107
6.9921 | -102.5465
6.9958 | -102.5802
6.9995 | -102.6186
7.0033 | -102.6523
7.007 | -102.6851
7.0107 | -102.7159
7.0144 | -102.7474
7.0181 | -102.7688
7.0218 | -102.7896
7.0256 | -102.803
7.0293 | -102.8101
7.033 | -102.8109
7.0367 | -102.8071
7.0404 | -102.7934
7.0442 | -102.7724
7.0479 | -102.7514
7.0516 | -102.733
7.0553 | -102.7143
7.059 | -102.6852
7.0628 | -102.6579
7.0665 | -102.6317
7.0702 | -102.6122
7.0739 | -102.587
7.0776 | -102.5628
7.0814 | -102.5445
7.0851 | -102.5255
7.0888 | -102.5096
7.0925 | -102.5004
7.0962 | -102.4961
7.1 | -102.503
7.1037 | -102.5214
7.1074 | -102.5447
7.1111 | -102.5697
7.1148 | -102.5961
7.1185 | -102.6309
7.1223 | -102.6647
7.126 | -102.6933
7.1297 | -102.7167
7.1334 | -102.736
7.1371 | -102.7494
7.1409 | -102.7561
7.1446 | -102.7577
7.1483 | -102.7607
7.152 | -102.7624
7.1557 | -102.7653
7.1595 | -102.77
7.1632 | -102.7832
7.1669 | -102.8062
7.1706 | -102.8319
7.1743 | -102.864
7.1781 | -102.9017
7.1818 | -102.9461
7.1855 | -102.9849
7.1892 | -103.0156
7.1929 | -103.0457
7.1967 | -103.0702
7.2004 | -103.0858
7.2041 | -103.0841
7.2078 | -103.0656
7.2115 | -103.0397
7.2152 | -103.0031
7.219 | -102.9623
7.2227 | -102.92
7.2264 | -102.8719
7.2301 | -102.8274
7.2338 | -102.7814
7.2376 | -102.7391
7.2413 | -102.7008
7.245 | -102.6654
7.2487 | -102.6364
7.2524 | -102.613
7.2562 | -102.5882
7.2599 | -102.5633
7.2636 | -102.5378
7.2673 | -102.5168
7.271 | -102.4949
7.2748 | -102.4763
7.2785 | -102.4626
7.2822 | -102.4526
7.2859 | -102.447
7.2896 | -102.4469
7.2934 | -102.4552
7.2971 | -102.4693
7.3008 | -102.4815
7.3045 | -102.5001
7.3082 | -102.5272
7.3119 | -102.5655
7.3157 | -102.6061
7.3194 | -102.6466
7.3231 | -102.6946
7.3268 | -102.7431
7.3305 | -102.7855
7.3343 | -102.8159
7.338 | -102.8402
7.3417 | -102.8679
7.3454 | -102.893
7.3491 | -102.9075
7.3529 | -102.9145
7.3566 | -102.9148
7.3603 | -102.9125
7.364 | -102.9096
7.3677 | -102.9044
7.3715 | -102.8959
7.3752 | -102.8902
7.3789 | -102.8905
7.3826 | -102.8898
7.3863 | -102.894
7.3901 | -102.9059
7.3938 | -102.9193
7.3975 | -102.9361
7.4012 | -102.9568
7.4049 | -102.9808
7.4086 | -103.0127
7.4124 | -103.0413
7.4161 | -103.0667
7.4198 | -103.0916
7.4235 | -103.1085
7.4272 | -103.1244
7.431 | -103.1337
7.4347 | -103.138
7.4384 | -103.1382
7.4421 | -103.1395
7.4458 | -103.1364
7.4496 | -103.125
7.4533 | -103.1087
7.457 | -103.1006
7.4607 | -103.0986
7.4644 | -103.1036
7.4682 | -103.11
7.4719 | -103.1172
7.4756 | -103.1289
7.4793 | -103.1491
7.483 | -103.1652
7.4867 | -103.1797
7.4905 | -103.1903
7.4942 | -103.1965
7.4979 | -103.201
7.5016 | -103.2026
7.5053 | -103.1992
7.5091 | -103.1835
7.5128 | -103.1731
7.5165 | -103.1647
7.5202 | -103.1493
7.5239 | -103.1417
7.5277 | -103.1369
7.5314 | -103.1349
7.5351 | -103.1413
7.5388 | -103.1425
7.5425 | -103.1406
7.5463 | -103.1409
7.55 | -103.1429
7.5537 | -103.1438
7.5574 | -103.1436
7.5611 | -103.147
7.5649 | -103.1513
7.5686 | -103.1538
7.5723 | -103.158
7.576 | -103.1644
7.5797 | -103.1707
7.5835 | -103.1772
7.5872 | -103.1844
7.5909 | -103.1839
7.5946 | -103.1798
7.5983 | -103.1785
7.602 | -103.1762
7.6058 | -103.1672
7.6095 | -103.1555
7.6132 | -103.1449
7.6169 | -103.1357
7.6206 | -103.1321
7.6244 | -103.133
7.6281 | -103.14
7.6318 | -103.1488
7.6355 | -103.1655
7.6392 | -103.1826
7.643 | -103.2051
7.6467 | -103.2332
7.6504 | -103.2638
7.6541 | -103.2987
7.6578 | -103.3389
7.6616 | -103.3748
7.6653 | -103.4049
7.669 | -103.4322
7.6727 | -103.4544
7.6764 | -103.4693
7.6801 | -103.4796
7.6839 | -103.4923
7.6876 | -103.499
7.6913 | -103.5005
7.695 | -103.5043
7.6987 | -103.5111
7.7025 | -103.5143
7.7062 | -103.5192
7.7099 | -103.529
7.7136 | -103.5313
7.7173 | -103.5352
7.7211 | -103.5366
7.7248 | -103.5343
7.7285 | -103.5332
7.7322 | -103.5351
7.7359 | -103.5327
7.7397 | -103.5266
7.7434 | -103.5211
7.7471 | -103.5081
7.7508 | -103.4945
7.7545 | -103.4761
7.7583 | -103.4555
7.762 | -103.4321
7.7657 | -103.4077
7.7694 | -103.3804
7.7731 | -103.3531
7.7768 | -103.3266
7.7806 | -103.3011
7.7843 | -103.2781
7.788 | -103.2577
7.7917 | -103.2365
7.7954 | -103.2149
7.7992 | -103.1869
7.8029 | -103.1606
7.8066 | -103.1282
7.8103 | -103.0884
7.814 | -103.047
7.8178 | -103.0074
7.8215 | -102.9662
7.8252 | -102.9179
7.8289 | -102.8774
7.8326 | -102.8413
7.8364 | -102.8075
7.8401 | -102.7761
7.8438 | -102.7494
7.8475 | -102.7269
7.8512 | -102.7117
7.855 | -102.7069
7.8587 | -102.7092
7.8624 | -102.7231
7.8661 | -102.7442
7.8698 | -102.7749
7.8735 | -102.8144
7.8773 | -102.86
7.881 | -102.9141
7.8847 | -102.9749
7.8884 | -103.0379
7.8921 | -103.1012
7.8959 | -103.1624
7.8996 | -103.2209
7.9033 | -103.2734
7.907 | -103.3188
7.9107 | -103.3581
7.9145 | -103.3847
7.9182 | -103.4012
7.9219 | -103.4084
7.9256 | -103.4087
7.9293 | -103.3944
7.9331 | -103.367
7.9368 | -103.3317
7.9405 | -103.2975
7.9442 | -103.2536
7.9479 | -103.2029
7.9516 | -103.1566
7.9554 | -103.1106
7.9591 | -103.0674
7.9628 | -103.034
7.9665 | -103.0139
7.9702 | -103.0008
7.974 | -103.0013
7.9777 | -103.0151
7.9814 | -103.038
7.9851 | -103.0703
7.9888 | -103.111
7.9926 | -103.1591
7.9963 | -103.2072
8 | -103.2553
8.0037 | -103.3068
8.0074 | -103.3604
8.0112 | -103.409
8.0149 | -103.4486
8.0186 | -103.4763
8.0223 | -103.4998
8.026 | -103.5152
8.0298 | -103.525
8.0335 | -103.5281
8.0372 | -103.5238
8.0409 | -103.51
8.0446 | -103.4936
8.0483 | -103.4813
8.0521 | -103.4683
8.0558 | -103.4563
8.0595 | -103.4418
8.0632 | -103.4305
8.0669 | -103.4236
8.0707 | -103.4139
8.0744 | -103.4093
8.0781 | -103.4073
8.0818 | -103.4033
8.0855 | -103.3977
8.0893 | -103.392
8.093 | -103.3892
8.0967 | -103.3856
8.1004 | -103.3844
8.1041 | -103.3816
8.1079 | -103.3728
8.1116 | -103.36
8.1153 | -103.3529
8.119 | -103.3491
8.1227 | -103.3387
8.1265 | -103.3244
8.1302 | -103.3132
8.1339 | -103.3108
8.1376 | -103.3114
8.1413 | -103.3113
8.145 | -103.3095
8.1488 | -103.3138
8.1525 | -103.3201
8.1562 | -103.3296
8.1599 | -103.3477
8.1636 | -103.3674
8.1674 | -103.3878
8.1711 | -103.412
8.1748 | -103.4433
8.1785 | -103.4792
8.1822 | -103.5144
8.186 | -103.5479
8.1897 | -103.5844
8.1934 | -103.6273
8.1971 | -103.6701
8.2008 | -103.7081
8.2046 | -103.7444
8.2083 | -103.7855
8.212 | -103.8297
8.2157 | -103.8606
8.2194 | -103.8956
8.2232 | -103.9274
8.2269 | -103.959
8.2306 | -103.9888
8.2343 | -104.0133
8.238 | -104.0242
8.2417 | -104.0286
8.2455 | -104.0278
8.2492 | -104.0262
8.2529 | -104.0234
8.2566 | -104.0068
8.2603 | -103.9849
8.2641 | -103.9613
8.2678 | -103.9376
8.2715 | -103.9066
8.2752 | -103.8722
8.2789 | -103.832
8.2827 | -103.7857
8.2864 | -103.7351
8.2901 | -103.6837
8.2938 | -103.6296
8.2975 | -103.571
8.3013 | -103.5108
8.305 | -103.4547
8.3087 | -103.4004
8.3124 | -103.3456
8.3161 | -103.2923
8.3199 | -103.2479
8.3236 | -103.2086
8.3273 | -103.1772
8.331 | -103.1463
8.3347 | -103.1207
8.3384 | -103.1063
8.3422 | -103.0928
8.3459 | -103.0865
8.3496 | -103.0852
8.3533 | -103.0909
8.357 | -103.1036
8.3608 | -103.1202
8.3645 | -103.1412
8.3682 | -103.1653
8.3719 | -103.191
8.3756 | -103.2184
8.3794 | -103.2539
8.3831 | -103.2928
8.3868 | -103.3284
8.3905 | -103.3663
8.3942 | -103.4091
8.398 | -103.4556
8.4017 | -103.5058
8.4054 | -103.553
8.4091 | -103.5981
8.4128 | -103.6458
8.4166 | -103.6885
8.4203 | -103.7271
8.424 | -103.7599
8.4277 | -103.787
8.4314 | -103.8114
8.4351 | -103.8284
8.4389 | -103.8332
8.4426 | -103.827
8.4463 | -103.8191
8.45 | -103.8001
8.4537 | -103.7758
8.4575 | -103.749
8.4612 | -103.7191
8.4649 | -103.6882
8.4686 | -103.6552
8.4723 | -103.6225
8.4761 | -103.5962
8.4798 | -103.5775
8.4835 | -103.5646
8.4872 | -103.5566
8.4909 | -103.5594
8.4947 | -103.5698
8.4984 | -103.5899
8.5021 | -103.6169
8.5058 | -103.6493
8.5095 | -103.6904
8.5132 | -103.7333
8.517 | -103.772
8.5207 | -103.8074
8.5244 | -103.8459
8.5281 | -103.8824
8.5318 | -103.9195
8.5356 | -103.9495
8.5393 | -103.9828
8.543 | -104.0124
8.5467 | -104.0387
8.5504 | -104.0606
8.5542 | -104.0731
8.5579 | -104.0831
8.5616 | -104.0935
8.5653 | -104.1006
8.569 | -104.1082
8.5728 | -104.1153
8.5765 | -104.119
8.5802 | -104.1132
8.5839 | -104.1036
8.5876 | -104.1027
8.5914 | -104.1017
8.5951 | -104.1007
8.5988 | -104.1009
8.6025 | -104.1007
8.6062 | -104.0993
8.6099 | -104.0995
8.6137 | -104.0919
8.6174 | -104.0744
8.6211 | -104.0587
8.6248 | -104.0402
8.6285 | -104.0199
8.6323 | -104.0022
8.636 | -103.9823
8.6397 | -103.964
8.6434 | -103.9513
8.6471 | -103.9419
8.6509 | -103.9289
8.6546 | -103.9189
8.6583 | -103.9124
8.662 | -103.9064
8.6657 | -103.9009
8.6695 | -103.8974
8.6732 | -103.8954
8.6769 | -103.8902
8.6806 | -103.889
8.6843 | -103.8912
8.6881 | -103.8893
8.6918 | -103.8863
8.6955 | -103.8864
8.6992 | -103.8931
8.7029 | -103.9043
8.7066 | -103.9188
8.7104 | -103.935
8.7141 | -103.9489
8.7178 | -103.9619
8.7215 | -103.9786
8.7252 | -104.0026
8.729 | -104.0288
8.7327 | -104.0549
8.7364 | -104.0832
8.7401 | -104.1172
8.7438 | -104.1551
8.7476 | -104.1983
8.7513 | -104.2339
8.755 | -104.2647
8.7587 | -104.2943
8.7624 | -104.3235
8.7662 | -104.3483
8.7699 | -104.3689
8.7736 | -104.3822
8.7773 | -104.3888
8.781 | -104.3936
8.7848 | -104.387
8.7885 | -104.3711
8.7922 | -104.3549
8.7959 | -104.3401
8.7996 | -104.3217
8.8033 | -104.2983
8.8071 | -104.2788
8.8108 | -104.257
8.8145 | -104.2337
8.8182 | -104.2186
8.8219 | -104.2048
8.8257 | -104.1891
8.8294 | -104.1755
8.8331 | -104.1627
8.8368 | -104.1517
8.8405 | -104.1425
8.8443 | -104.1337
8.848 | -104.1238
8.8517 | -104.1103
8.8554 | -104.0899
8.8591 | -104.0664
8.8629 | -104.0408
8.8666 | -104.0087
8.8703 | -103.9771
8.874 | -103.9484
8.8777 | -103.9177
8.8814 | -103.8894
8.8852 | -103.864
8.8889 | -103.8401
8.8926 | -103.8209
8.8963 | -103.8096
8.9 | -103.8025
8.9038 | -103.8027
8.9075 | -103.808
8.9112 | -103.8204
8.9149 | -103.8414
8.9186 | -103.864
8.9224 | -103.8855
8.9261 | -103.913
8.9298 | -103.9431
8.9335 | -103.9729
8.9372 | -103.9944
8.941 | -104.0165
8.9447 | -104.0343
8.9484 | -104.0453
8.9521 | -104.0516
8.9558 | -104.0542
8.9596 | -104.0508
8.9633 | -104.0468
8.967 | -104.0414
8.9707 | -104.0311
8.9744 | -104.0221
8.9781 | -104.0177
8.9819 | -104.016
8.9856 | -104.0208
8.9893 | -104.0285
8.993 | -104.0377
8.9967 | -104.0523
9.0005 | -104.0657
9.0042 | -104.0748
9.0079 | -104.0858
9.0116 | -104.0938
9.0153 | -104.1028
9.0191 | -104.1129
9.0228 | -104.1169
9.0265 | -104.1123
9.0302 | -104.1066
9.0339 | -104.096
9.0377 | -104.0856
9.0414 | -104.0741
9.0451 | -104.0623
9.0488 | -104.0483
9.0525 | -104.0314
9.0563 | -104.0093
9.06 | -103.988
9.0637 | -103.9735
9.0674 | -103.962
9.0711 | -103.9452
9.0748 | -103.9276
9.0786 | -103.9137
9.0823 | -103.902
9.086 | -103.8915
9.0897 | -103.8855
9.0934 | -103.882
9.0972 | -103.8812
9.1009 | -103.8844
9.1046 | -103.8859
9.1083 | -103.8957
9.112 | -103.913
9.1158 | -103.9276
9.1195 | -103.9573
9.1232 | -103.9778
9.1269 | -103.9999
9.1306 | -104.0233
9.1344 | -104.0443
9.1381 | -104.0692
9.1418 | -104.0937
9.1455 | -104.1182
9.1492 | -104.1361
9.153 | -104.1534
9.1567 | -104.1734
9.1604 | -104.1932
9.1641 | -104.2132
9.1678 | -104.2338
9.1715 | -104.2497
9.1753 | -104.2612
9.179 | -104.2747
9.1827 | -104.2917
9.1864 | -104.3133
9.1901 | -104.3357
9.1939 | -104.3608
9.1976 | -104.383
9.2013 | -104.4081
9.205 | -104.4329
9.2087 | -104.459
9.2125 | -104.4811
9.2162 | -104.5031
9.2199 | -104.522
9.2236 | -104.5363
9.2273 | -104.5464
9.2311 | -104.5564
9.2348 | -104.5637
9.2385 | -104.5679
9.2422 | -104.5631
9.2459 | -104.5512
9.2497 | -104.5365
9.2534 | -104.5191
9.2571 | -104.4984
9.2608 | -104.4785
9.2645 | -104.4535
9.2682 | -104.4248
9.272 | -104.3987
9.2757 | -104.3695
9.2794 | -104.338
9.2831 | -104.31
9.2868 | -104.2832
9.2906 | -104.2558
9.2943 | -104.231
9.298 | -104.21
9.3017 | -104.1899
9.3054 | -104.1724
9.3092 | -104.1591
9.3129 | -104.1486
9.3166 | -104.1416
9.3203 | -104.1379
9.324 | -104.1366
9.3278 | -104.1395
9.3315 | -104.145
9.3352 | -104.156
9.3389 | -104.1676
9.3426 | -104.1823
9.3464 | -104.2011
9.3501 | -104.2213
9.3538 | -104.2385
9.3575 | -104.2569
9.3612 | -104.2746
9.3649 | -104.2909
9.3687 | -104.3012
9.3724 | -104.3117
9.3761 | -104.3202
9.3798 | -104.3253
9.3835 | -104.3267
9.3873 | -104.3256
9.391 | -104.3152
9.3947 | -104.3012
9.3984 | -104.2875
9.4021 | -104.2731
9.4059 | -104.257
9.4096 | -104.2364
9.4133 | -104.2192
9.417 | -104.2029
9.4207 | -104.1851
9.4245 | -104.1738
9.4282 | -104.1655
9.4319 | -104.1596
9.4356 | -104.1615
9.4393 | -104.1656
9.443 | -104.1743
9.4468 | -104.187
9.4505 | -104.2018
9.4542 | -104.221
9.4579 | -104.2419
9.4616 | -104.2634
9.4654 | -104.2864
9.4691 | -104.3103
9.4728 | -104.3323
9.4765 | -104.354
9.4802 | -104.3752
9.484 | -104.3917
9.4877 | -104.4072
9.4914 | -104.4175
9.4951 | -104.4239
9.4988 | -104.4267
9.5026 | -104.4252
9.5063 | -104.4227
9.51 | -104.4145
9.5137 | -104.403
9.5174 | -104.3915
9.5212 | -104.3825
9.5249 | -104.3711
9.5286 | -104.3549
9.5323 | -104.34
9.536 | -104.3208
9.5397 | -104.3055
9.5435 | -104.2958
9.5472 | -104.2889
9.5509 | -104.2794
9.5546 | -104.2706
9.5583 | -104.2677
9.5621 | -104.2657
9.5658 | -104.2631
9.5695 | -104.2626
9.5732 | -104.2622
9.5769 | -104.2661
9.5807 | -104.2727
9.5844 | -104.2836
9.5881 | -104.2973
9.5918 | -104.3068
9.5955 | -104.3254
9.5993 | -104.3501
9.603 | -104.3769
9.6067 | -104.4037
9.6104 | -104.4321
9.6141 | -104.4611
9.6179 | -104.4875
9.6216 | -104.5149
9.6253 | -104.5351
9.629 | -104.5602
9.6327 | -104.5787
9.6364 | -104.5898
9.6402 | -104.6013
9.6439 | -104.6104
9.6476 | -104.6154
9.6513 | -104.6178
9.655 | -104.6157
9.6588 | -104.6185
9.6625 | -104.6153
9.6662 | -104.6074
9.6699 | -104.5985
9.6736 | -104.5873
9.6774 | -104.5758
9.6811 | -104.5596
9.6848 | -104.5401
9.6885 | -104.5208
9.6922 | -104.5018
9.696 | -104.4817
9.6997 | -104.4534
9.7034 | -104.4263
9.7071 | -104.4019
9.7108 | -104.3788
9.7146 | -104.3572
9.7183 | -104.3386
9.722 | -104.3214
9.7257 | -104.3077
9.7294 | -104.2981
9.7331 | -104.2951
9.7369 | -104.2932
9.7406 | -104.294
9.7443 | -104.2978
9.748 | -104.3091
9.7517 | -104.3218
9.7555 | -104.3376
9.7592 | -104.3517
9.7629 | -104.3637
9.7666 | -104.3779
9.7703 | -104.3867
9.7741 | -104.3966
9.7778 | -104.4054
9.7815 | -104.4072
9.7852 | -104.406
9.7889 | -104.4037
9.7927 | -104.3997
9.7964 | -104.3906
9.8001 | -104.3802
9.8038 | -104.3699
9.8075 | -104.3582
9.8113 | -104.351
9.815 | -104.3476
9.8187 | -104.3448
9.8224 | -104.3452
9.8261 | -104.3481
9.8298 | -104.3516
9.8336 | -104.3576
9.8373 | -104.3725
9.841 | -104.3934
9.8447 | -104.42
9.8484 | -104.446
9.8522 | -104.4717
9.8559 | -104.4941
9.8596 | -104.5148
9.8633 | -104.5307
9.867 | -104.5465
9.8708 | -104.5579
9.8745 | -104.5656
9.8782 | -104.569
9.8819 | -104.5695
9.8856 | -104.5693
9.8894 | -104.57
9.8931 | -104.5671
9.8968 | -104.5625
9.9005 | -104.5575
9.9042 | -104.5573
9.9079 | -104.5582
9.9117 | -104.5602
9.9154 | -104.562
9.9191 | -104.5631
9.9228 | -104.5609
9.9265 | -104.5589
9.9303 | -104.5615
9.934 | -104.565
9.9377 | -104.5671
9.9414 | -104.5663
9.9451 | -104.5651
9.9489 | -104.5625
9.9526 | -104.5616
9.9563 | -104.5593
9.96 | -104.5554
9.9637 | -104.5475
9.9675 | -104.5391
9.9712 | -104.5338
9.9749 | -104.5291
9.9786 | -104.525
9.9823 | -104.5165
9.9861 | -104.5103
9.9898 | -104.5065
9.9935 | -104.504
9.9972 | -104.5054
10.0009 | -104.5118
10.0046 | -104.5227
10.0084 | -104.5362
10.0121 | -104.5489
10.0158 | -104.5625
10.0195 | -104.5814
10.0232 | -104.5977
10.027 | -104.6116
10.0307 | -104.6235
10.0344 | -104.6356
10.0381 | -104.6506
10.0418 | -104.6659
10.0456 | -104.6796
10.0493 | -104.6886
10.053 | -104.6976
10.0567 | -104.7072
10.0604 | -104.7192
10.0642 | -104.7317
10.0679 | -104.7459
10.0716 | -104.7618
10.0753 | -104.7696
10.079 | -104.7786
10.0828 | -104.7928
10.0865 | -104.8078
10.0902 | -104.8201
10.0939 | -104.8315
10.0976 | -104.8392
10.1013 | -104.8406
10.1051 | -104.8377
10.1088 | -104.8329
10.1125 | -104.8249
10.1162 | -104.8084
10.1199 | -104.7865
10.1237 | -104.7612
10.1274 | -104.7346
10.1311 | -104.7025
10.1348 | -104.6701
10.1385 | -104.6391
10.1423 | -104.6047
10.146 | -104.568
10.1497 | -104.5339
10.1534 | -104.5054
10.1571 | -104.4803
10.1609 | -104.4593
10.1646 | -104.4361
10.1683 | -104.4205
10.172 | -104.4118
10.1757 | -104.4094
10.1795 | -104.4101
10.1832 | -104.4111
10.1869 | -104.4111
10.1906 | -104.4138
10.1943 | -104.4314
10.198 | -104.4424
10.2018 | -104.4524
10.2055 | -104.4663
10.2092 | -104.4783
10.2129 | -104.4863
10.2166 | -104.4928
10.2204 | -104.4987
10.2241 | -104.5054
10.2278 | -104.5105
10.2315 | -104.5083
10.2352 | -104.5011
10.239 | -104.4908
10.2427 | -104.4785
10.2464 | -104.4669
10.2501 | -104.4558
10.2538 | -104.4486
10.2576 | -104.444
10.2613 | -104.441
10.265 | -104.4443
10.2687 | -104.45
10.2724 | -104.457
10.2762 | -104.4668
10.2799 | -104.4792
10.2836 | -104.4947
10.2873 | -104.5128
10.291 | -104.535
10.2947 | -104.5592
10.2985 | -104.5816
10.3022 | -104.6053
10.3059 | -104.6294
10.3096 | -104.6497
10.3133 | -104.6758
10.3171 | -104.702
10.3208 | -104.724
10.3245 | -104.7425
10.3282 | -104.761
10.3319 | -104.7807
10.3357 | -104.8018
10.3394 | -104.8177
10.3431 | -104.8265
10.3468 | -104.8329
10.3505 | -104.8374
10.3543 | -104.8385
10.358 | -104.8329
10.3617 | -104.829
10.3654 | -104.8258
10.3691 | -104.8198
10.3728 | -104.8055
10.3766 | -104.79
10.3803 | -104.7737
10.384 | -104.7586
10.3877 | -104.7467
10.3914 | -104.7387
10.3952 | -104.7311
10.3989 | -104.7242
10.4026 | -104.7171
10.4063 | -104.7072
10.41 | -104.7003
10.4138 | -104.6955
10.4175 | -104.6944
10.4212 | -104.6976
10.4249 | -104.7062
10.4286 | -104.7148
10.4324 | -104.7249
10.4361 | -104.7345
10.4398 | -104.7455
10.4435 | -104.7634
10.4472 | -104.7839
10.451 | -104.8021
10.4547 | -104.8193
10.4584 | -104.8371
10.4621 | -104.8582
10.4658 | -104.8699
10.4695 | -104.8762
10.4733 | -104.8806
10.477 | -104.884
10.4807 | -104.8822
10.4844 | -104.8753
10.4881 | -104.8651
10.4919 | -104.8479
10.4956 | -104.8292
10.4993 | -104.8105
10.503 | -104.7884
10.5067 | -104.7658
10.5105 | -104.7447
10.5142 | -104.7228
10.5179 | -104.6994
10.5216 | -104.6774
10.5253 | -104.6573
10.5291 | -104.6405
10.5328 | -104.6247
10.5365 | -104.611
10.5402 | -104.5974
10.5439 | -104.5824
10.5477 | -104.5714
10.5514 | -104.5652
10.5551 | -104.5624
10.5588 | -104.5628
10.5625 | -104.5669
10.5662 | -104.5751
10.57 | -104.5827
10.5737 | -104.5862
10.5774 | -104.5904
10.5811 | -104.6005
10.5848 | -104.6119
10.5886 | -104.6226
10.5923 | -104.631
10.596 | -104.6326
10.5997 | -104.6319
10.6034 | -104.6322
10.6072 | -104.6284
10.6109 | -104.6208
10.6146 | -104.6137
10.6183 | -104.6062
10.622 | -104.5989
10.6258 | -104.5936
10.6295 | -104.5859
10.6332 | -104.5782
10.6369 | -104.5707
10.6406 | -104.5652
10.6444 | -104.5594
10.6481 | -104.5489
10.6518 | -104.5423
10.6555 | -104.5346
10.6592 | -104.5261
10.6629 | -104.5206
10.6667 | -104.5153
10.6704 | -104.5124
10.6741 | -104.5097
10.6778 | -104.5061
10.6815 | -104.5029
10.6853 | -104.5046
10.689 | -104.5074
10.6927 | -104.5105
10.6964 | -104.5176
10.7001 | -104.5228
10.7039 | -104.5225
10.7076 | -104.5261
10.7113 | -104.5327
10.715 | -104.5391
10.7187 | -104.5458
10.7225 | -104.555
10.7262 | -104.5688
10.7299 | -104.5768
10.7336 | -104.5842
10.7373 | -104.5889
10.7411 | -104.5942
10.7448 | -104.6021
10.7485 | -104.6104
10.7522 | -104.6182
10.7559 | -104.6293
10.7596 | -104.6387
10.7634 | -104.6441
10.7671 | -104.6514
10.7708 | -104.658
10.7745 | -104.6653
10.7782 | -104.6719
10.782 | -104.676
10.7857 | -104.6755
10.7894 | -104.6753
10.7931 | -104.679
10.7968 | -104.683
10.8006 | -104.6901
10.8043 | -104.6945
10.808 | -104.6965
10.8117 | -104.7026
10.8154 | -104.7085
10.8192 | -104.71
10.8229 | -104.7111
10.8266 | -104.7139
10.8303 | -104.7185
10.834 | -104.7276
10.8377 | -104.7347
10.8415 | -104.735
10.8452 | -104.7354
10.8489 | -104.738
10.8526 | -104.7424
10.8563 | -104.7485
10.8601 | -104.7567
10.8638 | -104.7623
10.8675 | -104.7669
10.8712 | -104.7705
10.8749 | -104.7752
10.8787 | -104.7836
10.8824 | -104.7921
10.8861 | -104.8016
10.8898 | -104.8115
10.8935 | -104.8237
10.8973 | -104.836
10.901 | -104.8488
10.9047 | -104.8628
10.9084 | -104.875
10.9121 | -104.8885
10.9159 | -104.8998
10.9196 | -104.9123
10.9233 | -104.9288
10.927 | -104.9443
10.9307 | -104.9582
10.9344 | -104.9701
10.9382 | -104.9835
10.9419 | -104.9925
10.9456 | -105.0042
10.9493 | -105.0121
10.953 | -105.0122
10.9568 | -105.0092
10.9605 | -105.0052
10.9642 | -105.0021
10.9679 | -104.9937
10.9716 | -104.9876
10.9754 | -104.9831
10.9791 | -104.9731
10.9828 | -104.9655
10.9865 | -104.9591
10.9902 | -104.9506
10.994 | -104.9423
10.9977 | -104.9367
11.0014 | -104.9368
11.0051 | -104.9386
11.0088 | -104.9392
11.0126 | -104.9359
11.0163 | -104.9375
11.02 | -104.9444
11.0237 | -104.9496
11.0274 | -104.9585
11.0311 | -104.9702
11.0349 | -104.9842
11.0386 | -104.9999
11.0423 | -105.0176
11.046 | -105.0348
11.0497 | -105.057
11.0535 | -105.0771
11.0572 | -105.0969
11.0609 | -105.1151
11.0646 | -105.1312
11.0683 | -105.1491
11.0721 | -105.167
11.0758 | -105.1815
11.0795 | -105.191
11.0832 | -105.1949
11.0869 | -105.1941
11.0907 | -105.1898
11.0944 | -105.1859
11.0981 | -105.1774
11.1018 | -105.1711
11.1055 | -105.1596
11.1093 | -105.1434
11.113 | -105.1294
11.1167 | -105.1136
11.1204 | -105.0955
11.1241 | -105.0747
11.1278 | -105.0547
11.1316 | -105.0328
11.1353 | -105.0163
11.139 | -105.0067
11.1427 | -104.9948
11.1464 | -104.9845
11.1502 | -104.9777
11.1539 | -104.9723
11.1576 | -104.9704
11.1613 | -104.9729
11.165 | -104.978
11.1688 | -104.9863
11.1725 | -105.0022
11.1762 | -105.0193
11.1799 | -105.0341
11.1836 | -105.0581
11.1874 | -105.0839
11.1911 | -105.1104
11.1948 | -105.1363
11.1985 | -105.163
11.2022 | -105.1934
11.206 | -105.2237
11.2097 | -105.2473
11.2134 | -105.2696
11.2171 | -105.2853
11.2208 | -105.302
11.2245 | -105.3132
11.2283 | -105.3173
11.232 | -105.3148
11.2357 | -105.3075
11.2394 | -105.2992
11.2431 | -105.2874
11.2469 | -105.2659
11.2506 | -105.2416
11.2543 | -105.2126
11.258 | -105.1779
11.2617 | -105.1424
11.2655 | -105.1004
11.2692 | -105.0642
11.2729 | -105.032
11.2766 | -104.9986
11.2803 | -104.9644
11.2841 | -104.9357
11.2878 | -104.9112
11.2915 | -104.8905
11.2952 | -104.8709
11.2989 | -104.8459
11.3026 | -104.8259
11.3064 | -104.813
11.3101 | -104.8017
11.3138 | -104.7943
11.3175 | -104.7897
11.3212 | -104.788
11.325 | -104.7872
11.3287 | -104.7896
11.3324 | -104.7936
11.3361 | -104.7952
11.3398 | -104.7984
11.3436 | -104.8065
11.3473 | -104.8113
11.351 | -104.8162
11.3547 | -104.8246
11.3584 | -104.8342
11.3622 | -104.8447
11.3659 | -104.8506
11.3696 | -104.8577
11.3733 | -104.8645
11.377 | -104.87
11.3808 | -104.8777
11.3845 | -104.8807
11.3882 | -104.8815
11.3919 | -104.8796
11.3956 | -104.8778
11.3993 | -104.8727
11.4031 | -104.8644
11.4068 | -104.8562
11.4105 | -104.8503
11.4142 | -104.8436
11.4179 | -104.8397
11.4217 | -104.838
11.4254 | -104.8357
11.4291 | -104.8355
11.4328 | -104.8399
11.4365 | -104.8417
11.4403 | -104.8407
11.444 | -104.8383
11.4477 | -104.8415
11.4514 | -104.8444
11.4551 | -104.8481
11.4589 | -104.8525
11.4626 | -104.8622
11.4663 | -104.8788
11.47 | -104.8938
11.4737 | -104.905
11.4775 | -104.9171
11.4812 | -104.9295
11.4849 | -104.9412
11.4886 | -104.9516
11.4923 | -104.96
11.496 | -104.9667
11.4998 | -104.9745
11.5035 | -104.9815
11.5072 | -104.9866
11.5109 | -104.9885
11.5146 | -104.9894
11.5184 | -104.9857
11.5221 | -104.978
11.5258 | -104.9682
11.5295 | -104.9556
11.5332 | -104.9434
11.537 | -104.9273
11.5407 | -104.9104
11.5444 | -104.8905
11.5481 | -104.8699
11.5518 | -104.8481
11.5556 | -104.825
11.5593 | -104.8057
11.563 | -104.7844
11.5667 | -104.763
11.5704 | -104.7461
11.5742 | -104.7309
11.5779 | -104.7164
11.5816 | -104.7018
11.5853 | -104.6883
11.589 | -104.6794
11.5927 | -104.674
11.5965 | -104.6698
11.6002 | -104.6708
11.6039 | -104.6789
11.6076 | -104.6895
11.6113 | -104.7046
11.6151 | -104.7199
11.6188 | -104.734
11.6225 | -104.7514
11.6262 | -104.7699
11.6299 | -104.7869
11.6337 | -104.8065
11.6374 | -104.826
11.6411 | -104.8468
11.6448 | -104.8682
11.6485 | -104.8883
11.6523 | -104.9091
11.656 | -104.9275
11.6597 | -104.9449
11.6634 | -104.9598
11.6671 | -104.9682
11.6709 | -104.9759
11.6746 | -104.9864
11.6783 | -104.9974
11.682 | -105.0051
11.6857 | -105.0112
11.6894 | -105.0157
11.6932 | -105.0202
11.6969 | -105.0201
11.7006 | -105.0192
11.7043 | -105.0185
11.708 | -105.0134
11.7118 | -105.0101
11.7155 | -105.0045
11.7192 | -104.9964
11.7229 | -104.9924
11.7266 | -104.9879
11.7304 | -104.9874
11.7341 | -104.9891
11.7378 | -104.9838
11.7415 | -104.9805
11.7452 | -104.978
11.749 | -104.9761
11.7527 | -104.9749
11.7564 | -104.9736
11.7601 | -104.9731
11.7638 | -104.9767
11.7675 | -104.9807
11.7713 | -104.9821
11.775 | -104.9839
11.7787 | -104.981
11.7824 | -104.9801
11.7861 | -104.9864
11.7899 | -104.9907
11.7936 | -104.9926
11.7973 | -104.9941
11.801 | -104.9995
11.8047 | -104.9986
11.8085 | -105.0008
11.8122 | -105.0044
11.8159 | -105.0055
11.8196 | -105.0033
11.8233 | -105.0007
11.8271 | -105.0006
11.8308 | -105.0012
11.8345 | -105.0007
11.8382 | -104.9983
11.8419 | -104.9946
11.8457 | -104.9837
11.8494 | -104.9697
11.8531 | -104.958
11.8568 | -104.9485
11.8605 | -104.9392
11.8642 | -104.928
11.868 | -104.9209
11.8717 | -104.9103
11.8754 | -104.8996
11.8791 | -104.8891
11.8828 | -104.878
11.8866 | -104.8699
11.8903 | -104.8619
11.894 | -104.8488
11.8977 | -104.8362
11.9014 | -104.8233
11.9052 | -104.8119
11.9089 | -104.8046
11.9126 | -104.7978
11.9163 | -104.7917
11.92 | -104.7882
11.9238 | -104.7845
11.9275 | -104.7806
11.9312 | -104.7786
11.9349 | -104.7771
11.9386 | -104.7753
11.9424 | -104.7762
11.9461 | -104.7805
11.9498 | -104.7899
11.9535 | -104.796
11.9572 | -104.8002
11.9609 | -104.81
11.9647 | -104.8246
11.9684 | -104.8379
11.9721 | -104.8492
11.9758 | -104.8628
11.9795 | -104.8809
11.9833 | -104.8983
11.987 | -104.9166
11.9907 | -104.9338
11.9944 | -104.9509
11.9981 | -104.967
12.0019 | -104.9808
12.0056 | -104.9934
12.0093 | -105.006
12.013 | -105.0246
12.0167 | -105.0417
12.0205 | -105.0594
12.0242 | -105.0802
12.0279 | -105.1011
12.0316 | -105.1185
12.0353 | -105.1372
12.0391 | -105.1591
12.0428 | -105.1783
12.0465 | -105.2018
12.0502 | -105.2255
12.0539 | -105.2478
12.0576 | -105.2723
12.0614 | -105.293
12.0651 | -105.316
12.0688 | -105.3391
12.0725 | -105.3577
12.0762 | -105.3733
12.08 | -105.3876
12.0837 | -105.3964
12.0874 | -105.4021
12.0911 | -105.4038
12.0948 | -105.4022
12.0986 | -105.3994
12.1023 | -105.3934
12.106 | -105.384
12.1097 | -105.3698
12.1134 | -105.3582
12.1172 | -105.3452
12.1209 | -105.3293
12.1246 | -105.3129
12.1283 | -105.3015
12.132 | -105.289
12.1358 | -105.2721
12.1395 | -105.253
12.1432 | -105.2414
12.1469 | -105.2332
12.1506 | -105.2268
12.1543 | -105.2228
12.1581 | -105.2196
12.1618 | -105.2174
12.1655 | -105.222
12.1692 | -105.2279
12.1729 | -105.2338
12.1767 | -105.242
12.1804 | -105.2514
12.1841 | -105.2596
12.1878 | -105.2639
12.1915 | -105.2711
12.1953 | -105.2807
12.199 | -105.288
12.2027 | -105.293
12.2064 | -105.3004
12.2101 | -105.3082
12.2139 | -105.3121
12.2176 | -105.3165
12.2213 | -105.3229
12.225 | -105.3303
12.2287 | -105.3386
12.2325 | -105.3413
12.2362 | -105.3439
12.2399 | -105.3462
12.2436 | -105.349
12.2473 | -105.3494
12.251 | -105.3487
12.2548 | -105.347
12.2585 | -105.3456
12.2622 | -105.3369
12.2659 | -105.3307
12.2696 | -105.3229
12.2734 | -105.3117
12.2771 | -105.3006
12.2808 | -105.2865
12.2845 | -105.2726
12.2882 | -105.2553
12.292 | -105.239
12.2957 | -105.2235
12.2994 | -105.203
12.3031 | -105.1795
12.3068 | -105.1614
12.3106 | -105.1456
12.3143 | -105.132
12.318 | -105.1144
12.3217 | -105.0973
12.3254 | -105.0826
12.3291 | -105.0712
12.3329 | -105.0624
12.3366 | -105.0558
12.3403 | -105.0562
12.344 | -105.0494
12.3477 | -105.0483
12.3515 | -105.047
12.3552 | -105.0455
12.3589 | -105.0466
12.3626 | -105.0462
12.3663 | -105.0471
12.3701 | -105.0481
12.3738 | -105.0508
12.3775 | -105.0529
12.3812 | -105.055
12.3849 | -105.0581
12.3887 | -105.0609
12.3924 | -105.0678
12.3961 | -105.0746
12.3998 | -105.0799
12.4035 | -105.0854
12.4073 | -105.0898
12.411 | -105.0971
12.4147 | -105.1085
12.4184 | -105.1178
12.4221 | -105.1273
12.4258 | -105.1377
12.4296 | -105.1489
12.4333 | -105.1601
12.437 | -105.1711
12.4407 | -105.188
12.4444 | -105.203
12.4482 | -105.2163
12.4519 | -105.2307
12.4556 | -105.2471
12.4593 | -105.2651
12.463 | -105.2873
12.4668 | -105.3105
12.4705 | -105.3317
12.4742 | -105.3552
12.4779 | -105.3828
12.4816 | -105.4096
12.4854 | -105.432
12.4891 | -105.4549
12.4928 | -105.4779
12.4965 | -105.4994
12.5002 | -105.5197
12.504 | -105.5389
12.5077 | -105.5549
12.5114 | -105.5703
12.5151 | -105.5838
12.5188 | -105.5977
12.5225 | -105.613
12.5263 | -105.6229
12.53 | -105.6268
12.5337 | -105.6294
12.5374 | -105.6311
12.5411 | -105.627
12.5449 | -105.6244
12.5486 | -105.6183
12.5523 | -105.6094
12.556 | -105.597
12.5597 | -105.5859
12.5635 | -105.5757
12.5672 | -105.5644
12.5709 | -105.5492
12.5746 | -105.5287
12.5783 | -105.5092
12.5821 | -105.491
12.5858 | -105.4738
12.5895 | -105.4588
12.5932 | -105.4422
12.5969 | -105.4249
12.6007 | -105.4107
12.6044 | -105.3995
12.6081 | -105.3899
12.6118 | -105.377
12.6155 | -105.3667
12.6192 | -105.3596
12.623 | -105.3508
12.6267 | -105.3427
12.6304 | -105.3346
12.6341 | -105.3303
12.6378 | -105.327
12.6416 | -105.3217
12.6453 | -105.3121
12.649 | -105.3081
12.6527 | -105.3042
12.6564 | -105.2976
12.6602 | -105.2922
12.6639 | -105.2838
12.6676 | -105.2753
12.6713 | -105.2673
12.675 | -105.258
12.6788 | -105.2486
12.6825 | -105.2363
12.6862 | -105.2217
12.6899 | -105.2047
12.6936 | -105.1885
12.6974 | -105.175
12.7011 | -105.1593
12.7048 | -105.1439
12.7085 | -105.131
12.7122 | -105.1215
12.7159 | -105.1149
12.7197 | -105.1086
12.7234 | -105.1027
12.7271 | -105.0977
12.7308 | -105.0926
12.7345 | -105.0881
12.7383 | -105.085
12.742 | -105.0825
12.7457 | -105.0808
12.7494 | -105.0816
12.7531 | -105.0855
12.7569 | -105.0929
12.7606 | -105.1024
12.7643 | -105.1098
12.768 | -105.1197
12.7717 | -105.132
12.7755 | -105.1421
12.7792 | -105.1551
12.7829 | -105.1706
12.7866 | -105.185
12.7903 | -105.2043
12.794 | -105.2266
12.7978 | -105.2485
12.8015 | -105.2699
12.8052 | -105.292
12.8089 | -105.3159
12.8126 | -105.3398
12.8164 | -105.3601
12.8201 | -105.3781
12.8238 | -105.3958
12.8275 | -105.4112
12.8312 | -105.421
12.835 | -105.4258
12.8387 | -105.4285
12.8424 | -105.4269
12.8461 | -105.424
12.8498 | -105.4175
12.8536 | -105.4094
12.8573 | -105.4005
12.861 | -105.3883
12.8647 | -105.3758
12.8684 | -105.3651
12.8722 | -105.3568
12.8759 | -105.3478
12.8796 | -105.3416
12.8833 | -105.3309
12.887 | -105.3253
12.8907 | -105.3239
12.8945 | -105.3204
12.8982 | -105.3148
12.9019 | -105.3122
12.9056 | -105.3091
12.9093 | -105.3089
12.9131 | -105.3119
12.9168 | -105.3162
12.9205 | -105.32
12.9242 | -105.3242
12.9279 | -105.3259
12.9317 | -105.3259
12.9354 | -105.3262
12.9391 | -105.3268
12.9428 | -105.3265
12.9465 | -105.3248
12.9503 | -105.3203
12.954 | -105.3128
12.9577 | -105.3042
12.9614 | -105.2993
12.9651 | -105.2868
12.9689 | -105.2739
12.9726 | -105.2597
12.9763 | -105.2411
12.98 | -105.2233
12.9837 | -105.204
12.9874 | -105.1838
12.9912 | -105.1631
12.9949 | -105.1378
12.9986 | -105.111
13.0023 | -105.0867
13.006 | -105.0648
13.0098 | -105.0468
13.0135 | -105.0288
13.0172 | -105.0108
13.0209 | -104.9975
13.0246 | -104.9854
13.0284 | -104.9754
13.0321 | -104.9685
13.0358 | -104.9614
13.0395 | -104.9555
13.0432 | -104.9516
13.047 | -104.9541
13.0507 | -104.9578
13.0544 | -104.9632
13.0581 | -104.971
13.0618 | -104.9825
13.0656 | -104.9951
13.0693 | -105.01
13.073 | -105.0298
13.0767 | -105.0505
13.0804 | -105.0737
13.0841 | -105.1034
13.0879 | -105.1346
13.0916 | -105.1635
13.0953 | -105.197
13.099 | -105.2316
13.1027 | -105.2627
13.1065 | -105.2931
13.1102 | -105.3246
13.1139 | -105.3575
13.1176 | -105.3929
13.1213 | -105.4254
13.1251 | -105.4502
13.1288 | -105.4717
13.1325 | -105.4923
13.1362 | -105.5104
13.1399 | -105.5258
13.1437 | -105.5348
13.1474 | -105.538
13.1511 | -105.5361
13.1548 | -105.533
13.1585 | -105.5222
13.1623 | -105.5084
13.166 | -105.4952
13.1697 | -105.4775
13.1734 | -105.456
13.1771 | -105.434
13.1808 | -105.4106
13.1846 | -105.3928
13.1883 | -105.3767
13.192 | -105.3575
13.1957 | -105.336
13.1994 | -105.3141
13.2032 | -105.2962
13.2069 | -105.2817
13.2106 | -105.2681
13.2143 | -105.2554
13.218 | -105.2477
13.2218 | -105.2414
13.2255 | -105.237
13.2292 | -105.2337
13.2329 | -105.2346
13.2366 | -105.2335
13.2404 | -105.2367
13.2441 | -105.2393
13.2478 | -105.2417
13.2515 | -105.2447
13.2552 | -105.2474
13.2589 | -105.2494
13.2627 | -105.2497
13.2664 | -105.2456
13.2701 | -105.2379
13.2738 | -105.2338
13.2775 | -105.229
13.2813 | -105.2249
13.285 | -105.2187
13.2887 | -105.2121
13.2924 | -105.2077
13.2961 | -105.2022
13.2999 | -105.1948
13.3036 | -105.184
13.3073 | -105.1689
13.311 | -105.1578
13.3147 | -105.1521
13.3185 | -105.1441
13.3222 | -105.1333
13.3259 | -105.1212
13.3296 | -105.1112
13.3333 | -105.1046
13.3371 | -105.1005
13.3408 | -105.0962
13.3445 | -105.0924
13.3482 | -105.0886
13.3519 | -105.0854
13.3556 | -105.0816
13.3594 | -105.0792
13.3631 | -105.0794
13.3668 | -105.0812
13.3705 | -105.0854
13.3742 | -105.0863
13.378 | -105.0874
13.3817 | -105.0905
13.3854 | -105.0973
13.3891 | -105.1072
13.3928 | -105.1172
13.3966 | -105.1255
13.4003 | -105.1341
13.404 | -105.144
13.4077 | -105.1494
13.4114 | -105.1573
13.4152 | -105.1659
13.4189 | -105.1738
13.4226 | -105.1805
13.4263 | -105.186
13.43 | -105.1933
13.4338 | -105.1964
13.4375 | -105.2024
13.4412 | -105.2083
13.4449 | -105.2123
13.4486 | -105.2165
13.4523 | -105.2194
13.4561 | -105.2211
13.4598 | -105.223
13.4635 | -105.2226
13.4672 | -105.2159
13.4709 | -105.211
13.4747 | -105.2048
13.4784 | -105.1958
13.4821 | -105.189
13.4858 | -105.1813
13.4895 | -105.1718
13.4933 | -105.1629
13.497 | -105.1547
13.5007 | -105.1474
13.5044 | -105.1388
13.5081 | -105.1311
13.5119 | -105.1282
13.5156 | -105.125
13.5193 | -105.1231
13.523 | -105.1239
13.5267 | -105.1227
13.5305 | -105.1224
13.5342 | -105.1238
13.5379 | -105.1297
13.5416 | -105.1381
13.5453 | -105.1484
13.549 | -105.1624
13.5528 | -105.1783
13.5565 | -105.1948
13.5602 | -105.2125
13.5639 | -105.2287
13.5676 | -105.2489
13.5714 | -105.2729
13.5751 | -105.2957
13.5788 | -105.3168
13.5825 | -105.3392
13.5862 | -105.3621
13.59 | -105.3854
13.5937 | -105.4066
13.5974 | -105.4266
13.6011 | -105.447
13.6048 | -105.4669
13.6086 | -105.4815
13.6123 | -105.4953
13.616 | -105.5095
13.6197 | -105.5209
13.6234 | -105.5275
13.6272 | -105.5294
13.6309 | -105.5285
13.6346 | -105.5281
13.6383 | -105.5239
13.642 | -105.5194
13.6457 | -105.5112
13.6495 | -105.4993
13.6532 | -105.4887
13.6569 | -105.478
13.6606 | -105.467
13.6643 | -105.4564
13.6681 | -105.4411
13.6718 | -105.4264
13.6755 | -105.4129
13.6792 | -105.4003
13.6829 | -105.389
13.6867 | -105.3836
13.6904 | -105.3768
13.6941 | -105.3705
13.6978 | -105.3664
13.7015 | -105.3625
13.7053 | -105.3583
13.709 | -105.3541
13.7127 | -105.3545
13.7164 | -105.3564
13.7201 | -105.3577
13.7238 | -105.3614
13.7276 | -105.3684
13.7313 | -105.3782
13.735 | -105.3876
13.7387 | -105.3973
13.7424 | -105.4072
13.7462 | -105.4167
13.7499 | -105.4291
13.7536 | -105.4429
13.7573 | -105.4585
13.761 | -105.4729
13.7648 | -105.485
13.7685 | -105.4977
13.7722 | -105.5085
13.7759 | -105.5183
13.7796 | -105.5231
13.7834 | -105.5303
13.7871 | -105.5371
13.7908 | -105.5462
13.7945 | -105.5543
13.7982 | -105.5581
13.802 | -105.5572
13.8057 | -105.5564
13.8094 | -105.5518
13.8131 | -105.5476
13.8168 | -105.5411
13.8205 | -105.5334
13.8243 | -105.5256
13.828 | -105.5171
13.8317 | -105.503
13.8354 | -105.4903
13.8391 | -105.4759
13.8429 | -105.4609
13.8466 | -105.4495
13.8503 | -105.4338
13.854 | -105.42
13.8577 | -105.4065
13.8615 | -105.3926
13.8652 | -105.3825
13.8689 | -105.3733
13.8726 | -105.3656
13.8763 | -105.3543
13.8801 | -105.3476
13.8838 | -105.3434
13.8875 | -105.3436
13.8912 | -105.3473
13.8949 | -105.3529
13.8987 | -105.3587
13.9024 | -105.3663
13.9061 | -105.3745
13.9098 | -105.3819
13.9135 | -105.3919
13.9172 | -105.4035
13.921 | -105.415
13.9247 | -105.4242
13.9284 | -105.4354
13.9321 | -105.4465
13.9358 | -105.4616
13.9396 | -105.4793
13.9433 | -105.4947
13.947 | -105.5072
13.9507 | -105.5204
13.9544 | -105.5347
13.9582 | -105.5431
13.9619 | -105.555
13.9656 | -105.5657
13.9693 | -105.5767
13.973 | -105.5903
13.9768 | -105.6038
13.9805 | -105.6177
13.9842 | -105.6251
13.9879 | -105.6296
13.9916 | -105.63
13.9954 | -105.6286
13.9991 | -105.6272
14.0028 | -105.6239
14.0065 | -105.6186
14.0102 | -105.6102
14.0139 | -105.5972
14.0177 | -105.5845
14.0214 | -105.5686
14.0251 | -105.5523
14.0288 | -105.5328
14.0325 | -105.5092
14.0363 | -105.484
14.04 | -105.46
14.0437 | -105.4364
14.0474 | -105.4116
14.0511 | -105.389
14.0549 | -105.3643
14.0586 | -105.3409
14.0623 | -105.3176
14.066 | -105.2962
14.0697 | -105.2788
14.0735 | -105.2654
14.0772 | -105.257
14.0809 | -105.2474
14.0846 | -105.2378
14.0883 | -105.232
14.0921 | -105.2288
14.0958 | -105.2301
14.0995 | -105.2344
14.1032 | -105.2415
14.1069 | -105.2516
14.1106 | -105.264
14.1144 | -105.2778
14.1181 | -105.2932
14.1218 | -105.3106
14.1255 | -105.3294
14.1292 | -105.3484
14.133 | -105.3708
14.1367 | -105.3895
14.1404 | -105.4073
14.1441 | -105.4293
14.1478 | -105.4536
14.1516 | -105.4782
14.1553 | -105.4982
14.159 | -105.5179
14.1627 | -105.5412
14.1664 | -105.5613
14.1702 | -105.5786
14.1739 | -105.5883
14.1776 | -105.5917
14.1813 | -105.5953
14.185 | -105.5989
14.1887 | -105.6018
14.1925 | -105.6033
14.1962 | -105.6021
14.1999 | -105.5949
14.2036 | -105.5873
14.2073 | -105.5804
14.2111 | -105.5729
14.2148 | -105.5625
14.2185 | -105.5518
14.2222 | -105.5363
14.2259 | -105.5247
14.2297 | -105.5135
14.2334 | -105.5044
14.2371 | -105.4964
14.2408 | -105.4861
14.2445 | -105.4746
14.2483 | -105.465
14.252 | -105.4548
14.2557 | -105.4476
14.2594 | -105.4427
14.2631 | -105.4384
14.2669 | -105.4357
14.2706 | -105.428
14.2743 | -105.4178
14.278 | -105.4089
14.2817 | -105.4019
14.2854 | -105.3973
14.2892 | -105.3926
14.2929 | -105.3878
14.2966 | -105.383
14.3003 | -105.3766
14.304 | -105.3714
14.3078 | -105.366
14.3115 | -105.3627
14.3152 | -105.3576
14.3189 | -105.3528
14.3226 | -105.3515
14.3264 | -105.3513
14.3301 | -105.3509
14.3338 | -105.3516
14.3375 | -105.3517
14.3412 | -105.3534
14.345 | -105.3582
14.3487 | -105.3637
14.3524 | -105.3686
14.3561 | -105.376
14.3598 | -105.3878
14.3636 | -105.3973
14.3673 | -105.4073
14.371 | -105.4231
14.3747 | -105.44
14.3784 | -105.4577
14.3821 | -105.4814
14.3859 | -105.5032
14.3896 | -105.5247
14.3933 | -105.5495
14.397 | -105.5743
14.4007 | -105.5999
14.4045 | -105.6279
14.4082 | -105.6533
14.4119 | -105.6763
14.4156 | -105.6972
14.4193 | -105.7197
14.4231 | -105.7429
14.4268 | -105.7677
14.4305 | -105.7854
14.4342 | -105.8
14.4379 | -105.8159
14.4417 | -105.832
14.4454 | -105.8478
14.4491 | -105.8561
14.4528 | -105.8621
14.4565 | -105.8679
14.4603 | -105.8716
14.464 | -105.871
14.4677 | -105.8688
14.4714 | -105.8678
14.4751 | -105.8679
14.4788 | -105.864
14.4826 | -105.8559
14.4863 | -105.8455
14.49 | -105.8366
14.4937 | -105.8253
14.4974 | -105.8156
14.5012 | -105.8038
14.5049 | -105.7896
14.5086 | -105.7787
14.5123 | -105.7668
14.516 | -105.7533
14.5198 | -105.7404
14.5235 | -105.733
14.5272 | -105.7257
14.5309 | -105.7148
14.5346 | -105.7078
14.5384 | -105.704
14.5421 | -105.7004
14.5458 | -105.6977
14.5495 | -105.6989
14.5532 | -105.6972
14.557 | -105.695
14.5607 | -105.6979
14.5644 | -105.6991
14.5681 | -105.6993
14.5718 | -105.703
14.5755 | -105.7071
14.5793 | -105.7104
14.583 | -105.7153
14.5867 | -105.7222
14.5904 | -105.7271
14.5941 | -105.7337
14.5979 | -105.7414
14.6016 | -105.7511
14.6053 | -105.7625
14.609 | -105.7735
14.6127 | -105.7859
14.6165 | -105.7991
14.6202 | -105.8111
14.6239 | -105.8254
14.6276 | -105.8397
14.6313 | -105.8514
14.6351 | -105.8616
14.6388 | -105.8731
14.6425 | -105.8833
14.6462 | -105.8888
14.6499 | -105.8908
14.6536 | -105.8922
14.6574 | -105.8972
14.6611 | -105.8971
14.6648 | -105.8925
14.6685 | -105.8845
14.6722 | -105.8773
14.676 | -105.8686
14.6797 | -105.855
14.6834 | -105.8398
14.6871 | -105.8229
14.6908 | -105.8062
14.6946 | -105.7888
14.6983 | -105.7734
14.702 | -105.7584
14.7057 | -105.7415
14.7094 | -105.7224
14.7132 | -105.7055
14.7169 | -105.6882
14.7206 | -105.6727
14.7243 | -105.659
14.728 | -105.6485
14.7318 | -105.6375
14.7355 | -105.6266
14.7392 | -105.6176
14.7429 | -105.6131
14.7466 | -105.6083
14.7503 | -105.6032
14.7541 | -105.6007
14.7578 | -105.6022
14.7615 | -105.6078
14.7652 | -105.6163
14.7689 | -105.6261
14.7727 | -105.636
14.7764 | -105.6495
14.7801 | -105.6626
14.7838 | -105.6755
14.7875 | -105.6886
14.7913 | -105.702
14.795 | -105.7187
14.7987 | -105.7352
14.8024 | -105.7495
14.8061 | -105.7637
14.8099 | -105.7773
14.8136 | -105.7884
14.8173 | -105.7999
14.821 | -105.8124
14.8247 | -105.8225
14.8285 | -105.8268
14.8322 | -105.8284
14.8359 | -105.8308
14.8396 | -105.83
14.8433 | -105.8273
14.847 | -105.823
14.8508 | -105.8154
14.8545 | -105.8065
14.8582 | -105.7955
14.8619 | -105.7842
14.8656 | -105.771
14.8694 | -105.7563
14.8731 | -105.7407
14.8768 | -105.7232
14.8805 | -105.7022
14.8842 | -105.6813
14.888 | -105.6596
14.8917 | -105.6385
14.8954 | -105.6158
14.8991 | -105.5977
14.9028 | -105.5792
14.9066 | -105.5591
14.9103 | -105.5366
14.914 | -105.52
14.9177 | -105.5066
14.9214 | -105.4945
14.9252 | -105.4805
14.9289 | -105.4696
14.9326 | -105.4569
14.9363 | -105.4444
14.94 | -105.4322
14.9437 | -105.4229
14.9475 | -105.4119
14.9512 | -105.3991
14.9549 | -105.3902
14.9586 | -105.381
14.9623 | -105.3737
14.9661 | -105.3659
14.9698 | -105.3573
14.9735 | -105.3462
14.9772 | -105.3354
14.9809 | -105.3276
14.9847 | -105.3158
14.9884 | -105.3048
14.9921 | -105.2982
14.9958 | -105.2931
14.9995 | -105.2871
15.0033 | -105.2804
15.007 | -105.274
15.0107 | -105.2695
15.0144 | -105.2663
15.0181 | -105.2641
15.0219 | -105.2706
15.0256 | -105.2716
15.0293 | -105.2737
15.033 | -105.2793
15.0367 | -105.2863
15.0404 | -105.2937
15.0442 | -105.3047
15.0479 | -105.3206
15.0516 | -105.3318
15.0553 | -105.345
15.059 | -105.3585
15.0628 | -105.373
15.0665 | -105.3907
15.0702 | -105.4073
15.0739 | -105.426
15.0776 | -105.4425
15.0814 | -105.4594
15.0851 | -105.4746
15.0888 | -105.4899
15.0925 | -105.5063
15.0962 | -105.5239
15.1 | -105.5407
15.1037 | -105.557
15.1074 | -105.5723
15.1111 | -105.5877
15.1148 | -105.5976
15.1186 | -105.6048
15.1223 | -105.6146
15.126 | -105.6247
15.1297 | -105.6292
15.1334 | -105.6328
15.1371 | -105.635
15.1409 | -105.6384
15.1446 | -105.6378
15.1483 | -105.6344
15.152 | -105.633
15.1557 | -105.6317
15.1595 | -105.6243
15.1632 | -105.6173
15.1669 | -105.613
15.1706 | -105.605
15.1743 | -105.5956
15.1781 | -105.5853
15.1818 | -105.575
15.1855 | -105.564
15.1892 | -105.5538
15.1929 | -105.5434
15.1967 | -105.5307
15.2004 | -105.5227
15.2041 | -105.5165
15.2078 | -105.508
15.2115 | -105.4974
15.2152 | -105.4882
15.219 | -105.4818
15.2227 | -105.4745
15.2264 | -105.4664
15.2301 | -105.4639
15.2338 | -105.4614
15.2376 | -105.4587
15.2413 | -105.4592
15.245 | -105.4575
15.2487 | -105.459
15.2524 | -105.4592
15.2562 | -105.4622
15.2599 | -105.4672
15.2636 | -105.4745
15.2673 | -105.4812
15.271 | -105.4899
15.2748 | -105.4981
15.2785 | -105.5085
15.2822 | -105.5199
15.2859 | -105.5273
15.2896 | -105.5357
15.2934 | -105.5467
15.2971 | -105.5595
15.3008 | -105.5703
15.3045 | -105.5809
15.3082 | -105.5901
15.3119 | -105.6008
15.3157 | -105.6093
15.3194 | -105.6124
15.3231 | -105.6156
15.3268 | -105.619
15.3305 | -105.6215
15.3343 | -105.6254
15.338 | -105.6298
15.3417 | -105.6299
15.3454 | -105.6273
15.3491 | -105.6286
15.3529 | -105.6305
15.3566 | -105.6294
15.3603 | -105.6246
15.364 | -105.6211
15.3677 | -105.6161
15.3715 | -105.6092
15.3752 | -105.6011
15.3789 | -105.595
15.3826 | -105.5897
15.3863 | -105.5823
15.3901 | -105.5734
15.3938 | -105.5682
15.3975 | -105.5619
15.4012 | -105.5546
15.4049 | -105.5491
15.4086 | -105.5416
15.4124 | -105.5352
15.4161 | -105.5296
15.4198 | -105.5211
15.4235 | -105.5092
15.4272 | -105.4999
15.431 | -105.4921
15.4347 | -105.4832
15.4384 | -105.4705
15.4421 | -105.4581
15.4458 | -105.4477
15.4496 | -105.4355
15.4533 | -105.4224
15.457 | -105.4112
15.4607 | -105.4003
15.4644 | -105.3881
15.4682 | -105.3755
15.4719 | -105.366
15.4756 | -105.3553
15.4793 | -105.3472
15.483 | -105.3407
15.4868 | -105.3351
15.4905 | -105.3308
15.4942 | -105.327
15.4979 | -105.3251
15.5016 | -105.3241
15.5053 | -105.3261
15.5091 | -105.3284
15.5128 | -105.3301
15.5165 | -105.3333
15.5202 | -105.3363
15.5239 | -105.3394
15.5277 | -105.341
15.5314 | -105.3438
15.5351 | -105.3513
15.5388 | -105.3612
15.5425 | -105.3689
15.5463 | -105.3752
15.55 | -105.3829
15.5537 | -105.3902
15.5574 | -105.3938
15.5611 | -105.3956
15.5649 | -105.4002
15.5686 | -105.407
15.5723 | -105.4124
15.576 | -105.4172
15.5797 | -105.4244
15.5834 | -105.4311
15.5872 | -105.4377
15.5909 | -105.4449
15.5946 | -105.4509
15.5983 | -105.4525
15.602 | -105.4556
15.6058 | -105.4599
15.6095 | -105.4625
15.6132 | -105.4648
15.6169 | -105.4653
15.6206 | -105.4635
15.6244 | -105.4641
15.6281 | -105.4649
15.6318 | -105.4671
15.6355 | -105.4681
15.6392 | -105.4684
15.643 | -105.4694
15.6467 | -105.4703
15.6504 | -105.4677
15.6541 | -105.4678
15.6578 | -105.4681
15.6616 | -105.469
15.6653 | -105.471
15.669 | -105.4693
15.6727 | -105.466
15.6764 | -105.4669
15.6801 | -105.4679
15.6839 | -105.4683
15.6876 | -105.4704
15.6913 | -105.4714
15.695 | -105.4694
15.6987 | -105.4689
15.7025 | -105.4692
15.7062 | -105.4672
15.7099 | -105.4654
15.7136 | -105.4671
15.7173 | -105.4685
15.7211 | -105.4654
15.7248 | -105.4644
15.7285 | -105.4649
15.7322 | -105.4668
15.7359 | -105.4682
15.7397 | -105.4707
15.7434 | -105.4743
15.7471 | -105.4763
15.7508 | -105.48
15.7545 | -105.479
15.7583 | -105.4745
15.762 | -105.4713
15.7657 | -105.4711
15.7694 | -105.4722
15.7731 | -105.472
15.7768 | -105.4697
15.7806 | -105.4705
15.7843 | -105.4698
15.788 | -105.4673
15.7917 | -105.4639
15.7954 | -105.4589
15.7992 | -105.4542
15.8029 | -105.4514
15.8066 | -105.4489
15.8103 | -105.446
15.814 | -105.4428
15.8178 | -105.4369
15.8215 | -105.4306
15.8252 | -105.4298
15.8289 | -105.4269
15.8326 | -105.4203
15.8364 | -105.4129
15.8401 | -105.4063
15.8438 | -105.3987
15.8475 | -105.3923
15.8512 | -105.3876
15.855 | -105.3811
15.8587 | -105.376
15.8624 | -105.3706
15.8661 | -105.3639
15.8698 | -105.3591
15.8735 | -105.3555
15.8773 | -105.3558
15.881 | -105.3528
15.8847 | -105.3466
15.8884 | -105.3407
15.8921 | -105.3334
15.8959 | -105.3283
15.8996 | -105.3265
15.9033 | -105.3214
15.907 | -105.3158
15.9107 | -105.3121
15.9145 | -105.3077
15.9182 | -105.3025
15.9219 | -105.2982
15.9256 | -105.2935
15.9293 | -105.2876
15.9331 | -105.2852
15.9368 | -105.2839
15.9405 | -105.2793
15.9442 | -105.2776
15.9479 | -105.2781
15.9517 | -105.2761
15.9554 | -105.2754
15.9591 | -105.2744
15.9628 | -105.274
15.9665 | -105.2732
15.9702 | -105.2737
15.974 | -105.2753
15.9777 | -105.2761
15.9814 | -105.2802
15.9851 | -105.2863
15.9888 | -105.2913
15.9926 | -105.2952
15.9963 | -105.298
16 | -105.3032
16.0037 | -105.3107
16.0074 | -105.3168
16.0112 | -105.3246
16.0149 | -105.334
16.0186 | -105.3424
16.0223 | -105.3521
16.026 | -105.3593
16.0298 | -105.3668
16.0335 | -105.3743
16.0372 | -105.3834
16.0409 | -105.3924
16.0446 | -105.4027
16.0483 | -105.4109
16.0521 | -105.4165
16.0558 | -105.4214
16.0595 | -105.4255
16.0632 | -105.4312
16.0669 | -105.4376
16.0707 | -105.4442
16.0744 | -105.4493
16.0781 | -105.4459
16.0818 | -105.4415
16.0855 | -105.4377
16.0893 | -105.4366
16.093 | -105.4354
16.0967 | -105.429
16.1004 | -105.4258
16.1041 | -105.4193
16.1079 | -105.4113
16.1116 | -105.4049
16.1153 | -105.4005
16.119 | -105.3955
16.1227 | -105.3898
16.1265 | -105.3843
16.1302 | -105.3751
16.1339 | -105.3661
16.1376 | -105.3585
16.1413 | -105.351
16.145 | -105.3457
16.1488 | -105.3409
16.1525 | -105.3369
16.1562 | -105.332
16.1599 | -105.3287
16.1636 | -105.3244
16.1674 | -105.3207
16.1711 | -105.3205
16.1748 | -105.3184
16.1785 | -105.3181
16.1822 | -105.3186
16.186 | -105.3168
16.1897 | -105.3151
16.1934 | -105.3157
16.1971 | -105.3158
16.2008 | -105.3151
16.2046 | -105.318
16.2083 | -105.3227
16.212 | -105.3267
16.2157 | -105.33
16.2194 | -105.3321
16.2232 | -105.3336
16.2269 | -105.3344
16.2306 | -105.3319
16.2343 | -105.3297
16.238 | -105.3281
16.2417 | -105.3274
16.2455 | -105.3231
16.2492 | -105.3155
16.2529 | -105.3095
16.2566 | -105.3052
16.2603 | -105.2976
16.2641 | -105.2901
16.2678 | -105.2825
16.2715 | -105.2733
16.2752 | -105.2655
16.2789 | -105.2581
16.2827 | -105.2495
16.2864 | -105.2406
16.2901 | -105.2341
16.2938 | -105.2272
16.2975 | -105.2199
16.3013 | -105.2133
16.305 | -105.2102
16.3087 | -105.2105
16.3124 | -105.213
16.3161 | -105.2153
16.3199 | -105.2178
16.3236 | -105.2231
16.3273 | -105.2275
16.331 | -105.232
16.3347 | -105.2388
16.3384 | -105.2483
16.3422 | -105.2569
16.3459 | -105.2676
16.3496 | -105.28
16.3533 | -105.2937
16.357 | -105.3065
16.3608 | -105.3184
16.3645 | -105.3315
16.3682 | -105.3455
16.3719 | -105.3583
16.3756 | -105.3704
16.3794 | -105.383
16.3831 | -105.3943
16.3868 | -105.4032
16.3905 | -105.4133
16.3942 | -105.4212
16.398 | -105.4284
16.4017 | -105.4344
16.4054 | -105.4386
16.4091 | -105.4442
16.4128 | -105.448
16.4166 | -105.4489
16.4203 | -105.4529
16.424 | -105.4576
16.4277 | -105.46
16.4314 | -105.4638
16.4351 | -105.465
16.4389 | -105.4629
16.4426 | -105.4598
16.4463 | -105.4563
16.45 | -105.4546
16.4537 | -105.4549
16.4575 | -105.4572
16.4612 | -105.4592
16.4649 | -105.462
16.4686 | -105.4632
16.4723 | -105.4641
16.4761 | -105.4657
16.4798 | -105.4706
16.4835 | -105.4757
16.4872 | -105.4793
16.4909 | -105.4816
16.4947 | -105.4853
16.4984 | -105.4884
16.5021 | -105.4938
16.5058 | -105.5013
16.5095 | -105.5063
16.5132 | -105.5115
16.517 | -105.5182
16.5207 | -105.5256
16.5244 | -105.532
16.5281 | -105.5399
16.5318 | -105.5473
16.5356 | -105.5546
16.5393 | -105.5621
16.543 | -105.5721
16.5467 | -105.5805
16.5504 | -105.5898
16.5542 | -105.5992
16.5579 | -105.6095
16.5616 | -105.6226
16.5653 | -105.6373
16.569 | -105.6494
16.5728 | -105.6586
16.5765 | -105.6678
16.5802 | -105.6733
16.5839 | -105.6782
16.5876 | -105.684
16.5914 | -105.6885
16.5951 | -105.6907
16.5988 | -105.692
16.6025 | -105.6938
16.6062 | -105.696
16.6099 | -105.6985
16.6137 | -105.6982
16.6174 | -105.698
16.6211 | -105.6984
16.6248 | -105.6971
16.6285 | -105.6939
16.6323 | -105.6892
16.636 | -105.6869
16.6397 | -105.6854
16.6434 | -105.6847
16.6471 | -105.6822
16.6509 | -105.6793
16.6546 | -105.6772
16.6583 | -105.6718
16.662 | -105.6654
16.6657 | -105.6613
16.6695 | -105.6581
16.6732 | -105.6578
16.6769 | -105.6546
16.6806 | -105.6502
16.6843 | -105.6475
16.6881 | -105.6457
16.6918 | -105.644
16.6955 | -105.6401
16.6992 | -105.639
16.7029 | -105.6368
16.7066 | -105.6329
16.7104 | -105.6299
16.7141 | -105.6275
16.7178 | -105.6253
16.7215 | -105.6218
16.7252 | -105.6184
16.729 | -105.6132
16.7327 | -105.6103
16.7364 | -105.6058
16.7401 | -105.5992
16.7438 | -105.593
16.7476 | -105.5871
16.7513 | -105.5798
16.755 | -105.5728
16.7587 | -105.5644
16.7624 | -105.5545
16.7662 | -105.5409
16.7699 | -105.5284
16.7736 | -105.5163
16.7773 | -105.5019
16.781 | -105.4855
16.7848 | -105.472
16.7885 | -105.4581
16.7922 | -105.4454
16.7959 | -105.4332
16.7996 | -105.42
16.8033 | -105.4064
16.8071 | -105.3957
16.8108 | -105.387
16.8145 | -105.3773
16.8182 | -105.3706
16.8219 | -105.3639
16.8257 | -105.3562
16.8294 | -105.3511
16.8331 | -105.3469
16.8368 | -105.3428
16.8405 | -105.3386
16.8443 | -105.3364
16.848 | -105.3364
16.8517 | -105.339
16.8554 | -105.3427
16.8591 | -105.3462
16.8629 | -105.3508
16.8666 | -105.3574
16.8703 | -105.3638
16.874 | -105.3712
16.8777 | -105.379
16.8815 | -105.3879
16.8852 | -105.3971
16.8889 | -105.4088
16.8926 | -105.4213
16.8963 | -105.4355
16.9 | -105.4528
16.9038 | -105.4719
16.9075 | -105.488
16.9112 | -105.5018
16.9149 | -105.5182
16.9186 | -105.5377
16.9224 | -105.5575
16.9261 | -105.5737
16.9298 | -105.5888
16.9335 | -105.6063
16.9372 | -105.6254
16.941 | -105.6439
16.9447 | -105.6635
16.9484 | -105.6828
16.9521 | -105.6995
16.9558 | -105.7129
16.9596 | -105.7241
16.9633 | -105.737
16.967 | -105.7465
16.9707 | -105.7556
16.9744 | -105.7623
16.9781 | -105.7661
16.9819 | -105.7702
16.9856 | -105.7711
16.9893 | -105.7728
16.993 | -105.7761
16.9967 | -105.7754
17.0005 | -105.7732
17.0042 | -105.7706
17.0079 | -105.7665
17.0116 | -105.7599
17.0153 | -105.7531
17.0191 | -105.747
17.0228 | -105.7394
17.0265 | -105.7288
17.0302 | -105.7192
17.0339 | -105.7087
17.0377 | -105.6986
17.0414 | -105.6912
17.0451 | -105.6834
17.0488 | -105.6748
17.0525 | -105.6675
17.0563 | -105.6604
17.06 | -105.6499
17.0637 | -105.6416
17.0674 | -105.6354
17.0711 | -105.6325
17.0748 | -105.6311
17.0786 | -105.6298
17.0823 | -105.6275
17.086 | -105.6246
17.0897 | -105.6221
17.0934 | -105.6231
17.0972 | -105.6248
17.1009 | -105.6263
17.1046 | -105.6291
17.1083 | -105.6334
17.112 | -105.6375
17.1158 | -105.6408
17.1195 | -105.6457
17.1232 | -105.6505
17.1269 | -105.6578
17.1306 | -105.6636
17.1344 | -105.6667
17.1381 | -105.6703
17.1418 | -105.6691
17.1455 | -105.665
17.1492 | -105.663
17.153 | -105.6621
17.1567 | -105.6592
17.1604 | -105.6537
17.1641 | -105.6452
17.1678 | -105.6286
17.1715 | -105.6184
17.1753 | -105.6053
17.179 | -105.5893
17.1827 | -105.5737
17.1864 | -105.5551
17.1901 | -105.5369
17.1939 | -105.5191
17.1976 | -105.501
17.2013 | -105.4815
17.205 | -105.4619
17.2087 | -105.444
17.2125 | -105.4253
17.2162 | -105.4048
17.2199 | -105.3841
17.2236 | -105.3658
17.2273 | -105.3451
17.2311 | -105.3256
17.2348 | -105.3048
17.2385 | -105.2855
17.2422 | -105.2688
17.2459 | -105.2559
17.2497 | -105.2436
17.2534 | -105.2347
17.2571 | -105.2278
17.2608 | -105.2214
17.2645 | -105.2166
17.2682 | -105.214
17.272 | -105.2119
17.2757 | -105.2116
17.2794 | -105.2122
17.2831 | -105.2144
17.2868 | -105.2171
17.2906 | -105.2225
17.2943 | -105.2298
17.298 | -105.2369
17.3017 | -105.2449
17.3054 | -105.2553
17.3092 | -105.2645
17.3129 | -105.2726
17.3166 | -105.2803
17.3203 | -105.2889
17.324 | -105.2983
17.3278 | -105.309
17.3315 | -105.3179
17.3352 | -105.3262
17.3389 | -105.3344
17.3426 | -105.3434
17.3464 | -105.3534
17.3501 | -105.3628
17.3538 | -105.3709
17.3575 | -105.3784
17.3612 | -105.3849
17.3649 | -105.3925
17.3687 | -105.4003
17.3724 | -105.4059
17.3761 | -105.4118
17.3798 | -105.422
17.3835 | -105.4298
17.3873 | -105.437
17.391 | -105.444
17.3947 | -105.4521
17.3984 | -105.4615
17.4021 | -105.4692
17.4059 | -105.4771
17.4096 | -105.4869
17.4133 | -105.4952
17.417 | -105.5033
17.4207 | -105.5114
17.4245 | -105.5185
17.4282 | -105.5264
17.4319 | -105.532
17.4356 | -105.5395
17.4393 | -105.5451
17.443 | -105.5482
17.4468 | -105.5525
17.4505 | -105.5567
17.4542 | -105.5587
17.4579 | -105.5593
17.4616 | -105.5594
17.4654 | -105.5581
17.4691 | -105.5544
17.4728 | -105.5498
17.4765 | -105.5412
17.4802 | -105.5328
17.484 | -105.5239
17.4877 | -105.5117
17.4914 | -105.4983
17.4951 | -105.4806
17.4988 | -105.4639
17.5026 | -105.4451
17.5063 | -105.4253
17.51 | -105.4051
17.5137 | -105.3832
17.5174 | -105.3607
17.5212 | -105.3388
17.5249 | -105.3156
17.5286 | -105.294
17.5323 | -105.2718
17.536 | -105.2503
17.5397 | -105.2293
17.5435 | -105.2093
17.5472 | -105.19
17.5509 | -105.1723
17.5546 | -105.1545
17.5583 | -105.1395
17.5621 | -105.1245
17.5658 | -105.1092
17.5695 | -105.0958
17.5732 | -105.0833
17.5769 | -105.069
17.5807 | -105.057
17.5844 | -105.0457
17.5881 | -105.0338
17.5918 | -105.0232
17.5955 | -105.014
17.5993 | -105.0032
17.603 | -104.9946
17.6067 | -104.9876
17.6104 | -104.9794
17.6141 | -104.9682
17.6179 | -104.9578
17.6216 | -104.9515
17.6253 | -104.9457
17.629 | -104.9381
17.6327 | -104.9318
17.6364 | -104.9271
17.6402 | -104.9237
17.6439 | -104.9197
17.6476 | -104.9151
17.6513 | -104.9117
17.655 | -104.9096
17.6588 | -104.9087
17.6625 | -104.9056
17.6662 | -104.9059
17.6699 | -104.9109
17.6736 | -104.9145
17.6774 | -104.9163
17.6811 | -104.9205
17.6848 | -104.9239
17.6885 | -104.9259
17.6922 | -104.93
17.696 | -104.9337
17.6997 | -104.9393
17.7034 | -104.9465
17.7071 | -104.952
17.7108 | -104.9552
17.7146 | -104.9593
17.7183 | -104.9642
17.722 | -104.9682
17.7257 | -104.9723
17.7294 | -104.9775
17.7331 | -104.9821
17.7369 | -104.9872
17.7406 | -104.9932
17.7443 | -104.9961
17.748 | -105.001
17.7517 | -105.0074
17.7555 | -105.0132
17.7592 | -105.0194
17.7629 | -105.027
17.7666 | -105.0358
17.7703 | -105.0441
17.7741 | -105.0541
17.7778 | -105.0635
17.7815 | -105.0691
17.7852 | -105.0765
17.7889 | -105.0859
17.7927 | -105.0945
17.7964 | -105.102
17.8001 | -105.1101
17.8038 | -105.1181
17.8075 | -105.1274
17.8113 | -105.1356
17.815 | -105.1446
17.8187 | -105.1536
17.8224 | -105.1617
17.8261 | -105.1666
17.8298 | -105.1729
17.8336 | -105.1803
17.8373 | -105.1866
17.841 | -105.1924
17.8447 | -105.1968
17.8484 | -105.2019
17.8522 | -105.21
17.8559 | -105.2155
17.8596 | -105.218
17.8633 | -105.2222
17.867 | -105.2283
17.8708 | -105.2326
17.8745 | -105.2337
17.8782 | -105.2359
17.8819 | -105.2377
17.8856 | -105.2399
17.8894 | -105.2417
17.8931 | -105.2429
17.8968 | -105.2423
17.9005 | -105.2422
17.9042 | -105.2415
17.9079 | -105.2385
17.9117 | -105.234
17.9154 | -105.231
17.9191 | -105.2257
17.9228 | -105.2189
17.9265 | -105.2137
17.9303 | -105.2069
17.934 | -105.1994
17.9377 | -105.1932
17.9414 | -105.1877
17.9451 | -105.1811
17.9489 | -105.1747
17.9526 | -105.1669
17.9563 | -105.1633
17.96 | -105.1561
17.9637 | -105.15
17.9675 | -105.1477
17.9712 | -105.1448
17.9749 | -105.1434
17.9786 | -105.1428
17.9823 | -105.1434
17.9861 | -105.1446
17.9898 | -105.1456
17.9935 | -105.1458
17.9972 | -105.1452
18.0009 | -105.1493
18.0046 | -105.1536
18.0084 | -105.1555
18.0121 | -105.1591
18.0158 | -105.1629
18.0195 | -105.1651
18.0232 | -105.1671
18.027 | -105.1703
18.0307 | -105.1737
18.0344 | -105.1805
18.0381 | -105.1911
18.0418 | -105.1994
18.0456 | -105.2051
18.0493 | -105.2133
18.053 | -105.223
18.0567 | -105.2333
18.0604 | -105.2436
18.0642 | -105.2536
18.0679 | -105.2632
18.0716 | -105.2725
18.0753 | -105.2802
18.079 | -105.2881
18.0828 | -105.3007
18.0865 | -105.3111
18.0902 | -105.3205
18.0939 | -105.3299
18.0976 | -105.3386
18.1013 | -105.3454
18.1051 | -105.3523
18.1088 | -105.3582
18.1125 | -105.363
18.1162 | -105.3697
18.1199 | -105.3747
18.1237 | -105.3782
18.1274 | -105.3794
18.1311 | -105.3801
18.1348 | -105.3815
18.1385 | -105.3821
18.1423 | -105.381
18.146 | -105.3803
18.1497 | -105.3796
18.1534 | -105.376
18.1571 | -105.3714
18.1609 | -105.3692
18.1646 | -105.3687
18.1683 | -105.3662
18.172 | -105.3641
18.1757 | -105.3618
18.1795 | -105.3587
18.1832 | -105.3581
18.1869 | -105.3583
18.1906 | -105.3573
18.1943 | -105.3546
18.198 | -105.3539
18.2018 | -105.3538
18.2055 | -105.3522
18.2092 | -105.3505
18.2129 | -105.3484
18.2166 | -105.3497
18.2204 | -105.3503
18.2241 | -105.3456
18.2278 | -105.3437
18.2315 | -105.3406
18.2352 | -105.3355
18.239 | -105.3266
18.2427 | -105.3165
18.2464 | -105.3093
18.2501 | -105.3031
18.2538 | -105.2935
18.2576 | -105.2813
18.2613 | -105.2705
18.265 | -105.258
18.2687 | -105.2458
18.2724 | -105.2332
18.2762 | -105.2194
18.2799 | -105.2055
18.2836 | -105.1924
18.2873 | -105.1766
18.291 | -105.16
18.2947 | -105.1401
18.2985 | -105.1194
18.3022 | -105.1023
18.3059 | -105.086
18.3096 | -105.0671
18.3133 | -105.0499
18.3171 | -105.034
18.3208 | -105.0175
18.3245 | -105.0011
18.3282 | -104.9868
18.3319 | -104.9734
18.3357 | -104.9631
18.3394 | -104.9525
18.3431 | -104.9404
18.3468 | -104.9322
18.3505 | -104.9224
18.3543 | -104.913
18.358 | -104.903
18.3617 | -104.8935
18.3654 | -104.885
18.3691 | -104.8779
18.3728 | -104.8738
18.3766 | -104.8689
18.3803 | -104.8649
18.384 | -104.8605
18.3877 | -104.8577
18.3914 | -104.8544
18.3952 | -104.8497
18.3989 | -104.8466
18.4026 | -104.8448
18.4063 | -104.8427
18.41 | -104.8392
18.4138 | -104.8369
18.4175 | -104.8324
18.4212 | -104.8258
18.4249 | -104.8183
18.4286 | -104.8133
18.4324 | -104.809
18.4361 | -104.8039
18.4398 | -104.7988
18.4435 | -104.7921
18.4472 | -104.7853
18.451 | -104.7775
18.4547 | -104.7717
18.4584 | -104.7653
18.4621 | -104.7578
18.4658 | -104.7511
18.4695 | -104.7452
18.4733 | -104.738
18.477 | -104.7295
18.4807 | -104.7225
18.4844 | -104.7184
18.4881 | -104.7121
18.4919 | -104.7063
18.4956 | -104.7035
18.4993 | -104.7029
18.503 | -104.7011
18.5067 | -104.6975
18.5105 | -104.694
18.5142 | -104.6913
18.5179 | -104.6932
18.5216 | -104.6956
18.5253 | -104.6945
18.5291 | -104.6986
18.5328 | -104.7037
18.5365 | -104.7085
18.5402 | -104.7139
18.5439 | -104.7197
18.5477 | -104.7231
18.5514 | -104.7295
18.5551 | -104.7336
18.5588 | -104.7387
18.5625 | -104.7459
18.5662 | -104.7513
18.57 | -104.7544
18.5737 | -104.7586
18.5774 | -104.7604
18.5811 | -104.7616
18.5848 | -104.7646
18.5886 | -104.767
18.5923 | -104.7673
18.596 | -104.7676
18.5997 | -104.7677
18.6034 | -104.767
18.6072 | -104.7629
18.6109 | -104.755
18.6146 | -104.7491
18.6183 | -104.7399
18.622 | -104.7316
18.6258 | -104.7252
18.6295 | -104.716
18.6332 | -104.7045
18.6369 | -104.6932
18.6406 | -104.6824
18.6444 | -104.6706
18.6481 | -104.661
18.6518 | -104.6501
18.6555 | -104.6387
18.6592 | -104.626
18.6629 | -104.6117
18.6667 | -104.597
18.6704 | -104.5846
18.6741 | -104.5703
18.6778 | -104.5559
18.6815 | -104.5436
18.6853 | -104.5295
18.689 | -104.5177
18.6927 | -104.506
18.6964 | -104.4921
18.7001 | -104.4794
18.7039 | -104.4644
18.7076 | -104.4498
18.7113 | -104.4362
18.715 | -104.4224
18.7187 | -104.406
18.7225 | -104.3876
18.7262 | -104.3709
18.7299 | -104.3527
18.7336 | -104.333
18.7373 | -104.3135
18.7411 | -104.2952
18.7448 | -104.277
18.7485 | -104.258
18.7522 | -104.2388
18.7559 | -104.2191
18.7596 | -104.1987
18.7634 | -104.1753
18.7671 | -104.1512
18.7708 | -104.1319
18.7745 | -104.1146
18.7782 | -104.0962
18.782 | -104.0785
18.7857 | -104.0599
18.7894 | -104.0432
18.7931 | -104.0272
18.7968 | -104.0095
18.8006 | -103.9904
18.8043 | -103.9737
18.808 | -103.9576
18.8117 | -103.9436
18.8154 | -103.9322
18.8192 | -103.9198
18.8229 | -103.9076
18.8266 | -103.8968
18.8303 | -103.8831
18.834 | -103.8704
18.8377 | -103.8594
18.8415 | -103.8458
18.8452 | -103.8342
18.8489 | -103.8188
18.8526 | -103.8047
18.8563 | -103.7928
18.8601 | -103.7782
18.8638 | -103.7639
18.8675 | -103.7509
18.8712 | -103.7391
18.8749 | -103.7272
18.8787 | -103.7157
18.8824 | -103.7025
18.8861 | -103.6918
18.8898 | -103.6815
18.8935 | -103.668
18.8973 | -103.652
18.901 | -103.636
18.9047 | -103.6213
18.9084 | -103.6092
18.9121 | -103.5963
18.9159 | -103.5835
18.9196 | -103.5719
18.9233 | -103.5605
18.927 | -103.5488
18.9307 | -103.5381
18.9344 | -103.5263
18.9382 | -103.5108
18.9419 | -103.4953
18.9456 | -103.4825
18.9493 | -103.4677
18.953 | -103.4542
18.9568 | -103.4355
18.9605 | -103.4174
18.9642 | -103.4017
18.9679 | -103.3859
18.9716 | -103.3714
18.9754 | -103.3553
18.9791 | -103.3386
18.9828 | -103.3226
18.9865 | -103.3075
18.9902 | -103.2945
18.994 | -103.2809
18.9977 | -103.2678
19.0014 | -103.2544
19.0051 | -103.2412
19.0088 | -103.2277
19.0126 | -103.2142
19.0163 | -103.2017
19.02 | -103.1908
19.0237 | -103.1817
19.0274 | -103.1715
19.0311 | -103.1638
19.0349 | -103.1571
19.0386 | -103.1487
19.0423 | -103.1403
19.046 | -103.1346
19.0497 | -103.1293
19.0535 | -103.1233
19.0572 | -103.1184
19.0609 | -103.1128
19.0646 | -103.1077
19.0683 | -103.1034
19.0721 | -103.1008
19.0758 | -103.0983
19.0795 | -103.0949
19.0832 | -103.0914
19.0869 | -103.0892
19.0907 | -103.0844
19.0944 | -103.0814
19.0981 | -103.0805
19.1018 | -103.0816
19.1055 | -103.0824
19.1093 | -103.0842
19.113 | -103.0877
19.1167 | -103.0902
19.1204 | -103.097
19.1241 | -103.1026
19.1278 | -103.1076
19.1316 | -103.1145
19.1353 | -103.1224
19.139 | -103.1296
19.1427 | -103.1376
19.1464 | -103.1497
19.1502 | -103.1631
19.1539 | -103.1776
19.1576 | -103.1924
19.1613 | -103.207
19.165 | -103.2242
19.1688 | -103.2457
19.1725 | -103.2669
19.1762 | -103.2875
19.1799 | -103.3066
19.1836 | -103.3274
19.1874 | -103.3488
19.1911 | -103.3718
19.1948 | -103.3944
19.1985 | -103.4171
19.2022 | -103.4398
19.206 | -103.4636
19.2097 | -103.4879
19.2134 | -103.5127
19.2171 | -103.5346
19.2208 | -103.5571
19.2245 | -103.5785
19.2283 | -103.6032
19.232 | -103.6253
19.2357 | -103.644
19.2394 | -103.6623
19.2431 | -103.6817
19.2469 | -103.7
19.2506 | -103.7164
19.2543 | -103.7314
19.258 | -103.7457
19.2617 | -103.7591
19.2655 | -103.771
19.2692 | -103.7818
19.2729 | -103.7939
19.2766 | -103.8058
19.2803 | -103.8154
19.2841 | -103.8232
19.2878 | -103.8299
19.2915 | -103.8354
19.2952 | -103.8408
19.2989 | -103.8456
19.3026 | -103.8487
19.3064 | -103.85
19.3101 | -103.851
19.3138 | -103.8482
19.3175 | -103.849
19.3212 | -103.8429
19.325 | -103.8375
19.3287 | -103.834
19.3324 | -103.8316
19.3361 | -103.8302
19.3398 | -103.8285
19.3436 | -103.825
19.3473 | -103.8201
19.351 | -103.8158
19.3547 | -103.8122
19.3584 | -103.8081
19.3622 | -103.8028
19.3659 | -103.7979
19.3696 | -103.7928
19.3733 | -103.7869
19.377 | -103.7823
19.3808 | -103.7779
19.3845 | -103.7747
19.3882 | -103.7728
19.3919 | -103.771
19.3956 | -103.7683
19.3993 | -103.7654
19.4031 | -103.7644
19.4068 | -103.7656
19.4105 | -103.7642
19.4142 | -103.7612
19.4179 | -103.7607
19.4217 | -103.7619
19.4254 | -103.7618
19.4291 | -103.7613
19.4328 | -103.7604
19.4365 | -103.7612
19.4403 | -103.7636
19.444 | -103.7643
19.4477 | -103.7649
19.4514 | -103.7627
19.4551 | -103.7612
19.4589 | -103.7589
19.4626 | -103.7575
19.4663 | -103.7557
19.47 | -103.7553
19.4737 | -103.7576
19.4775 | -103.7577
19.4812 | -103.7577
19.4849 | -103.7591
19.4886 | -103.7597
19.4923 | -103.7609
19.496 | -103.7631
19.4998 | -103.7646
19.5035 | -103.7642
19.5072 | -103.765
19.5109 | -103.7678
19.5146 | -103.7685
19.5184 | -103.769
19.5221 | -103.773
19.5258 | -103.7769
19.5295 | -103.779
19.5332 | -103.7812
19.537 | -103.7825
19.5407 | -103.7833
19.5444 | -103.7836
19.5481 | -103.7838
19.5518 | -103.784
19.5556 | -103.7864
19.5593 | -103.789
19.563 | -103.7926
19.5667 | -103.7947
19.5704 | -103.7947
19.5742 | -103.7975
19.5779 | -103.8017
19.5816 | -103.8075
19.5853 | -103.8125
19.589 | -103.8158
19.5927 | -103.8188
19.5965 | -103.8216
19.6002 | -103.832
19.6039 | -103.8418
19.6076 | -103.8502
19.6113 | -103.8574
19.6151 | -103.8627
19.6188 | -103.8676
19.6225 | -103.8758
19.6262 | -103.8839
19.6299 | -103.8931
19.6337 | -103.9021
19.6374 | -103.9108
19.6411 | -103.9196
19.6448 | -103.9306
19.6485 | -103.9403
19.6523 | -103.9485
19.656 | -103.9549
19.6597 | -103.9618
19.6634 | -103.968
19.6671 | -103.9731
19.6709 | -103.9782
19.6746 | -103.9828
19.6783 | -103.9864
19.682 | -103.9862
19.6857 | -103.986
19.6894 | -103.9882
19.6932 | -103.9938
19.6969 | -103.995
19.7006 | -103.9933
19.7043 | -103.9937
19.708 | -103.9933
19.7118 | -103.993
19.7155 | -103.9904
19.7192 | -103.9873
19.7229 | -103.9851
19.7266 | -103.9841
19.7304 | -103.9853
19.7341 | -103.9876
19.7378 | -103.9905
19.7415 | -103.9898
19.7452 | -103.9899
19.749 | -103.9911
19.7527 | -103.9898
19.7564 | -103.9893
19.7601 | -103.99
19.7638 | -103.9908
19.7675 | -103.992
19.7713 | -103.9929
19.775 | -103.9937
19.7787 | -103.9954
19.7824 | -103.9982
19.7861 | -104.0009
19.7899 | -104.0042
19.7936 | -104.0108
19.7973 | -104.0165
19.801 | -104.0205
19.8047 | -104.0254
19.8085 | -104.0324
19.8122 | -104.0396
19.8159 | -104.0469
19.8196 | -104.0563
19.8233 | -104.0673
19.8271 | -104.0767
19.8308 | -104.0872
19.8345 | -104.0957
19.8382 | -104.1043
19.8419 | -104.1126
19.8457 | -104.1246
19.8494 | -104.1402
19.8531 | -104.1535
19.8568 | -104.165
19.8605 | -104.1754
19.8642 | -104.1842
19.868 | -104.1956
19.8717 | -104.2092
19.8754 | -104.2223
19.8791 | -104.234
19.8828 | -104.243
19.8866 | -104.2525
19.8903 | -104.2618
19.894 | -104.2709
19.8977 | -104.2813
19.9014 | -104.2926
19.9052 | -104.3
19.9089 | -104.3066
19.9126 | -104.3105
19.9163 | -104.3141
19.92 | -104.3188
19.9238 | -104.3221
19.9275 | -104.3239
19.9312 | -104.3253
19.9349 | -104.3249
19.9386 | -104.3219
19.9424 | -104.3176
19.9461 | -104.3112
19.9498 | -104.3032
19.9535 | -104.2942
19.9572 | -104.2863
19.9609 | -104.2796
19.9647 | -104.2714
19.9684 | -104.2614
19.9721 | -104.2486
19.9758 | -104.2332
19.9795 | -104.2202
19.9833 | -104.2074
19.987 | -104.1944
19.9907 | -104.1835
19.9944 | -104.1724
19.9981 | -104.1619
20.0019 | -104.1508
20.0056 | -104.1375
20.0093 | -104.1246
20.013 | -104.1136
20.0167 | -104.1021
20.0205 | -104.091
20.0242 | -104.0839
20.0279 | -104.0753
20.0316 | -104.0665
20.0353 | -104.0605
20.0391 | -104.0561
20.0428 | -104.0517
20.0465 | -104.0473
20.0502 | -104.0441
20.0539 | -104.0437
20.0576 | -104.0438
20.0614 | -104.0453
20.0651 | -104.0474
20.0688 | -104.0494
20.0725 | -104.0517
20.0762 | -104.0545
20.08 | -104.0596
20.0837 | -104.0664
20.0874 | -104.0733
20.0911 | -104.0811
20.0948 | -104.0888
20.0986 | -104.0967
20.1023 | -104.1043
20.106 | -104.1129
20.1097 | -104.1221
20.1134 | -104.129
20.1172 | -104.1352
20.1209 | -104.144
20.1246 | -104.153
20.1283 | -104.1601
20.132 | -104.1677
20.1358 | -104.1724
20.1395 | -104.1751
20.1432 | -104.1789
20.1469 | -104.1822
20.1506 | -104.1843
20.1543 | -104.1847
20.1581 | -104.1823
20.1618 | -104.1782
20.1655 | -104.1754
20.1692 | -104.1713
20.1729 | -104.164
20.1767 | -104.156
20.1804 | -104.1475
20.1841 | -104.1357
20.1878 | -104.1213
20.1915 | -104.1092
20.1953 | -104.0994
20.199 | -104.0882
20.2027 | -104.0749
20.2064 | -104.0603
20.2101 | -104.0448
20.2139 | -104.0286
20.2176 | -104.0126
20.2213 | -103.9982
20.225 | -103.9845
20.2287 | -103.9694
20.2325 | -103.957
20.2362 | -103.9446
20.2399 | -103.9305
20.2436 | -103.9127
20.2473 | -103.8968
20.251 | -103.883
20.2548 | -103.8691
20.2585 | -103.8552
20.2622 | -103.8411
20.2659 | -103.8259
20.2696 | -103.8106
20.2734 | -103.7975
20.2771 | -103.782
20.2808 | -103.7678
20.2845 | -103.7563
20.2882 | -103.7434
20.292 | -103.7317
20.2957 | -103.7243
20.2994 | -103.7169
20.3031 | -103.708
20.3068 | -103.6991
20.3106 | -103.6869
20.3143 | -103.6763
20.318 | -103.6676
20.3217 | -103.6577
20.3254 | -103.6487
20.3291 | -103.6397
20.3329 | -103.631
20.3366 | -103.6216
20.3403 | -103.6103
20.344 | -103.6013
20.3477 | -103.5933
20.3515 | -103.5854
20.3552 | -103.5792
20.3589 | -103.5718
20.3626 | -103.5623
20.3663 | -103.5533
20.3701 | -103.5452
20.3738 | -103.5386
20.3775 | -103.531
20.3812 | -103.5244
20.3849 | -103.5172
20.3887 | -103.5183
20.3924 | -103.5149
20.3961 | -103.5113
20.3998 | -103.5096
20.4035 | -103.5082
20.4073 | -103.5071
20.411 | -103.5054
20.4147 | -103.5057
20.4184 | -103.5068
20.4221 | -103.5065
20.4258 | -103.5082
20.4296 | -103.5135
20.4333 | -103.5176
20.437 | -103.5218
20.4407 | -103.5269
20.4444 | -103.5312
20.4482 | -103.538
20.4519 | -103.5428
20.4556 | -103.5467
20.4593 | -103.5513
20.463 | -103.5571
20.4668 | -103.563
20.4705 | -103.5679
20.4742 | -103.5729
20.4779 | -103.5766
20.4816 | -103.5773
20.4854 | -103.5778
20.4891 | -103.5777
20.4928 | -103.5794
20.4965 | -103.5794
20.5002 | -103.578
20.504 | -103.5762
20.5077 | -103.5738
20.5114 | -103.5685
20.5151 | -103.5652
20.5188 | -103.5618
20.5225 | -103.5546
20.5263 | -103.5473
20.53 | -103.5417
20.5337 | -103.5363
20.5374 | -103.5274
20.5411 | -103.5191
20.5449 | -103.5109
20.5486 | -103.504
20.5523 | -103.4974
20.556 | -103.4904
20.5597 | -103.4852
20.5635 | -103.4777
20.5672 | -103.4704
20.5709 | -103.4641
20.5746 | -103.458
20.5783 | -103.454
20.5821 | -103.4515
20.5858 | -103.4439
20.5895 | -103.4375
20.5932 | -103.4339
20.5969 | -103.4307
20.6007 | -103.4301
20.6044 | -103.4311
20.6081 | -103.4322
20.6118 | -103.4345
20.6155 | -103.4356
20.6192 | -103.4384
20.623 | -103.4419
20.6267 | -103.4475
20.6304 | -103.4531
20.6341 | -103.455
20.6378 | -103.4577
20.6416 | -103.463
20.6453 | -103.4654
20.649 | -103.4711
20.6527 | -103.4792
20.6564 | -103.4862
20.6602 | -103.4918
20.6639 | -103.4969
20.6676 | -103.5019
20.6713 | -103.5054
20.675 | -103.5098
20.6788 | -103.5136
20.6825 | -103.515
20.6862 | -103.5173
20.6899 | -103.5172
20.6936 | -103.5153
20.6974 | -103.5143
20.7011 | -103.5149
20.7048 | -103.5136
20.7085 | -103.5087
20.7122 | -103.5032
20.7159 | -103.4972
20.7197 | -103.4888
20.7234 | -103.4814
20.7271 | -103.4696
20.7308 | -103.4578
20.7345 | -103.4451
20.7383 | -103.4314
20.742 | -103.4166
20.7457 | -103.4029
20.7494 | -103.3877
20.7531 | -103.3713
20.7569 | -103.3554
20.7606 | -103.3378
20.7643 | -103.3215
20.768 | -103.3055
20.7717 | -103.2894
20.7755 | -103.2743
20.7792 | -103.2566
20.7829 | -103.2401
20.7866 | -103.2235
20.7903 | -103.207
20.794 | -103.1911
20.7978 | -103.1748
20.8015 | -103.1576
20.8052 | -103.142
20.8089 | -103.1259
20.8126 | -103.1119
20.8164 | -103.1008
20.8201 | -103.0912
20.8238 | -103.0816
20.8275 | -103.0731
20.8312 | -103.0654
20.835 | -103.0564
20.8387 | -103.048
20.8424 | -103.0388
20.8461 | -103.0322
20.8498 | -103.027
20.8536 | -103.0218
20.8573 | -103.019
20.861 | -103.0181
20.8647 | -103.0171
20.8684 | -103.0152
20.8722 | -103.0145
20.8759 | -103.0148
20.8796 | -103.0197
20.8833 | -103.0239
20.887 | -103.0277
20.8907 | -103.0338
20.8945 | -103.0401
20.8982 | -103.044
20.9019 | -103.0482
20.9056 | -103.0535
20.9093 | -103.0592
20.9131 | -103.064
20.9168 | -103.0692
20.9205 | -103.0744
20.9242 | -103.0808
20.9279 | -103.0865
20.9317 | -103.096
20.9354 | -103.1036
20.9391 | -103.1113
20.9428 | -103.1207
20.9465 | -103.1275
20.9503 | -103.1328
20.954 | -103.1375
20.9577 | -103.1428
20.9614 | -103.1473
20.9651 | -103.1514
20.9689 | -103.157
20.9726 | -103.1624
20.9763 | -103.1669
20.98 | -103.1719
20.9837 | -103.1758
20.9874 | -103.1806
20.9912 | -103.1848
20.9949 | -103.1894
20.9986 | -103.1943
21.0023 | -103.1999
21.006 | -103.2047
21.0098 | -103.2092
21.0135 | -103.2141
21.0172 | -103.2206
21.0209 | -103.2243
21.0246 | -103.233
21.0284 | -103.2397
21.0321 | -103.2455
21.0358 | -103.2519
21.0395 | -103.259
21.0432 | -103.2671
21.047 | -103.2737
21.0507 | -103.281
21.0544 | -103.2888
21.0581 | -103.2974
21.0618 | -103.305
21.0656 | -103.3145
21.0693 | -103.3253
21.073 | -103.3364
21.0767 | -103.3494
21.0804 | -103.3594
21.0841 | -103.37
21.0879 | -103.3819
21.0916 | -103.3952
21.0953 | -103.4104
21.099 | -103.4266
21.1027 | -103.4424
21.1065 | -103.4572
21.1102 | -103.4714
21.1139 | -103.4857
21.1176 | -103.5003
21.1213 | -103.5144
21.1251 | -103.5281
21.1288 | -103.5436
21.1325 | -103.562
21.1362 | -103.5796
21.1399 | -103.597
21.1437 | -103.6112
21.1474 | -103.6254
21.1511 | -103.6412
21.1548 | -103.6557
21.1585 | -103.6732
21.1623 | -103.6914
21.166 | -103.7089
21.1697 | -103.7253
21.1734 | -103.74
21.1771 | -103.7518
21.1808 | -103.7655
21.1846 | -103.7814
21.1883 | -103.794
21.192 | -103.8067
21.1957 | -103.8216
21.1994 | -103.8333
21.2032 | -103.8435
21.2069 | -103.8553
21.2106 | -103.8667
21.2143 | -103.8778
21.218 | -103.8873
21.2218 | -103.8951
21.2255 | -103.9023
21.2292 | -103.9099
21.2329 | -103.9149
21.2366 | -103.9197
21.2404 | -103.9221
21.2441 | -103.9264
21.2478 | -103.9312
21.2515 | -103.9347
21.2552 | -103.9387
21.2589 | -103.9414
21.2627 | -103.9431
21.2664 | -103.945
21.2701 | -103.9453
21.2738 | -103.9445
21.2775 | -103.9426
21.2813 | -103.9394
21.285 | -103.9375
21.2887 | -103.9361
21.2924 | -103.9355
21.2961 | -103.9353
21.2999 | -103.9331
21.3036 | -103.9313
21.3073 | -103.9316
21.311 | -103.931
21.3147 | -103.9304
21.3185 | -103.9284
21.3222 | -103.9267
21.3259 | -103.9225
21.3296 | -103.921
21.3333 | -103.9217
21.3371 | -103.922
21.3408 | -103.9201
21.3445 | -103.919
21.3482 | -103.9199
21.3519 | -103.9211
21.3556 | -103.9221
21.3594 | -103.9228
21.3631 | -103.9218
21.3668 | -103.9205
21.3705 | -103.9208
21.3742 | -103.9201
21.378 | -103.9187
21.3817 | -103.9198
21.3854 | -103.9195
21.3891 | -103.9177
21.3928 | -103.9162
21.3966 | -103.914
21.4003 | -103.9105
21.404 | -103.9075
21.4077 | -103.9043
21.4114 | -103.901
21.4152 | -103.8964
21.4189 | -103.8902
21.4226 | -103.8845
21.4263 | -103.8787
21.43 | -103.8723
21.4338 | -103.8639
21.4375 | -103.8557
21.4412 | -103.8493
21.4449 | -103.8422
21.4486 | -103.8327
21.4523 | -103.8228
21.4561 | -103.813
21.4598 | -103.7984
21.4635 | -103.7891
21.4672 | -103.7794
21.4709 | -103.7707
21.4747 | -103.7634
21.4784 | -103.7562
21.4821 | -103.7498
21.4858 | -103.7426
21.4895 | -103.7347
21.4933 | -103.7273
21.497 | -103.7182
21.5007 | -103.7098
21.5044 | -103.7043
21.5081 | -103.6993
21.5119 | -103.6931
21.5156 | -103.6856
21.5193 | -103.6788
21.523 | -103.6726
21.5267 | -103.668
21.5305 | -103.6632
21.5342 | -103.657
21.5379 | -103.652
21.5416 | -103.6479
21.5453 | -103.6435
21.549 | -103.6384
21.5528 | -103.6333
21.5565 | -103.6303
21.5602 | -103.6284
21.5639 | -103.6255
21.5676 | -103.6221
21.5714 | -103.6194
21.5751 | -103.6157
21.5788 | -103.6098
21.5825 | -103.6072
21.5862 | -103.6037
21.59 | -103.6002
21.5937 | -103.5963
21.5974 | -103.5905
21.6011 | -103.5832
21.6048 | -103.5746
21.6086 | -103.5675
21.6123 | -103.5613
21.616 | -103.5552
21.6197 | -103.5472
21.6234 | -103.5417
21.6272 | -103.5356
21.6309 | -103.5295
21.6346 | -103.5199
21.6383 | -103.5095
21.642 | -103.4999
21.6457 | -103.4885
21.6495 | -103.4754
21.6532 | -103.4625
21.6569 | -103.4493
21.6606 | -103.4367
21.6643 | -103.4253
21.6681 | -103.4135
21.6718 | -103.3991
21.6755 | -103.3845
21.6792 | -103.3705
21.6829 | -103.3565
21.6867 | -103.3428
21.6904 | -103.3292
21.6941 | -103.3168
21.6978 | -103.3027
21.7015 | -103.2877
21.7053 | -103.2722
21.709 | -103.2564
21.7127 | -103.2404
21.7164 | -103.2257
21.7201 | -103.2113
21.7238 | -103.1967
21.7276 | -103.1816
21.7313 | -103.1698
21.735 | -103.1573
21.7387 | -103.1455
21.7424 | -103.1358
21.7462 | -103.125
21.7499 | -103.1134
21.7536 | -103.1033
21.7573 | -103.0945
21.761 | -103.0873
21.7648 | -103.081
21.7685 | -103.0759
21.7722 | -103.0687
21.7759 | -103.0594
21.7796 | -103.0486
21.7834 | -103.0384
21.7871 | -103.0295
21.7908 | -103.0216
21.7945 | -103.0156
21.7982 | -103.0095
21.802 | -103.0058
21.8057 | -103.0026
21.8094 | -102.9988
21.8131 | -102.9937
21.8168 | -102.9855
21.8205 | -102.9797
21.8243 | -102.9748
21.828 | -102.9688
21.8317 | -102.9628
21.8354 | -102.9571
21.8391 | -102.9526
21.8429 | -102.9479
21.8466 | -102.9416
21.8503 | -102.9354
21.854 | -102.9288
21.8577 | -102.9212
21.8615 | -102.9148
21.8652 | -102.9079
21.8689 | -102.9006
21.8726 | -102.8951
21.8763 | -102.8902
21.8801 | -102.8857
21.8838 | -102.8812
21.8875 | -102.8779
21.8912 | -102.8711
21.8949 | -102.8658
21.8987 | -102.8609
21.9024 | -102.856
21.9061 | -102.8512
21.9098 | -102.8455
21.9135 | -102.8431
21.9172 | -102.8386
21.921 | -102.8341
21.9247 | -102.831
21.9284 | -102.8294
21.9321 | -102.8269
21.9358 | -102.8247
21.9396 | -102.8233
21.9433 | -102.824
21.947 | -102.8231
21.9507 | -102.8227
21.9544 | -102.8231
21.9582 | -102.8249
21.9619 | -102.827
21.9656 | -102.8301
21.9693 | -102.8329
21.973 | -102.835
21.9768 | -102.8392
21.9805 | -102.8455
21.9842 | -102.8502
21.9879 | -102.8537
21.9916 | -102.8584
21.9954 | -102.8622
21.9991 | -102.8672
22.0028 | -102.871
22.0065 | -102.8747
22.0102 | -102.88
22.0139 | -102.886
22.0177 | -102.8923
22.0214 | -102.897
22.0251 | -102.9029
22.0288 | -102.9086
22.0325 | -102.9148
22.0363 | -102.9212
22.04 | -102.9267
22.0437 | -102.9296
22.0474 | -102.9316
22.0511 | -102.9331
22.0549 | -102.9339
22.0586 | -102.9373
22.0623 | -102.9422
22.066 | -102.9463
22.0697 | -102.9475
22.0735 | -102.9484
22.0772 | -102.9493
22.0809 | -102.9504
22.0846 | -102.9514
22.0883 | -102.9529
22.0921 | -102.9582
22.0958 | -102.9629
22.0995 | -102.9667
22.1032 | -102.9687
22.1069 | -102.9698
22.1106 | -102.9696
22.1144 | -102.9677
22.1181 | -102.9657
22.1218 | -102.9627
22.1255 | -102.9619
22.1292 | -102.9633
22.133 | -102.9608
22.1367 | -102.9589
22.1404 | -102.9576
22.1441 | -102.9556
22.1478 | -102.9526
22.1516 | -102.9472
22.1553 | -102.9433
22.159 | -102.9381
22.1627 | -102.9323
22.1664 | -102.9286
22.1702 | -102.923
22.1739 | -102.9169
22.1776 | -102.9111
22.1813 | -102.9054
22.185 | -102.8986
22.1887 | -102.892
22.1925 | -102.8841
22.1962 | -102.8756
22.1999 | -102.8692
22.2036 | -102.8638
22.2073 | -102.8551
22.2111 | -102.847
22.2148 | -102.8395
22.2185 | -102.8316
22.2222 | -102.8252
22.2259 | -102.8156
22.2297 | -102.8062
22.2334 | -102.7974
22.2371 | -102.7899
22.2408 | -102.783
22.2445 | -102.776
22.2483 | -102.7695
22.252 | -102.7634
22.2557 | -102.7561
22.2594 | -102.7475
22.2631 | -102.739
22.2669 | -102.7328
22.2706 | -102.7268
22.2743 | -102.722
22.278 | -102.7152
22.2817 | -102.709
22.2854 | -102.7033
22.2892 | -102.699
22.2929 | -102.694
22.2966 | -102.6879
22.3003 | -102.6846
22.304 | -102.6825
22.3078 | -102.6799
22.3115 | -102.6768
22.3152 | -102.6727
22.3189 | -102.6697
22.3226 | -102.6671
22.3264 | -102.6637
22.3301 | -102.6584
22.3338 | -102.6545
22.3375 | -102.6535
22.3412 | -102.6525
22.345 | -102.6494
22.3487 | -102.6454
22.3524 | -102.6429
22.3561 | -102.6422
22.3598 | -102.6409
22.3636 | -102.6414
22.3673 | -102.6434
22.371 | -102.6445
22.3747 | -102.6428
22.3784 | -102.6405
22.3821 | -102.6374
22.3859 | -102.6359
22.3896 | -102.6331
22.3933 | -102.6319
22.397 | -102.6313
22.4007 | -102.6314
22.4045 | -102.6323
22.4082 | -102.631
22.4119 | -102.6295
22.4156 | -102.6274
22.4193 | -102.6243
22.4231 | -102.6229
22.4268 | -102.623
22.4305 | -102.623
22.4342 | -102.6229
22.4379 | -102.6222
22.4417 | -102.6206
22.4454 | -102.6179
22.4491 | -102.6139
22.4528 | -102.606
22.4565 | -102.5984
22.4603 | -102.591
22.464 | -102.5818
22.4677 | -102.5735
22.4714 | -102.5646
22.4751 | -102.5552
22.4788 | -102.5434
22.4826 | -102.5305
22.4863 | -102.5162
22.49 | -102.4965
22.4937 | -102.4765
22.4974 | -102.4586
22.5012 | -102.4395
22.5049 | -102.4201
22.5086 | -102.3998
22.5123 | -102.3791
22.516 | -102.3558
22.5198 | -102.3324
22.5235 | -102.3096
22.5272 | -102.2851
22.5309 | -102.2583
22.5346 | -102.2248
22.5384 | -102.2008
22.5421 | -102.1748
22.5458 | -102.1475
22.5495 | -102.1198
22.5532 | -102.0924
22.557 | -102.0629
22.5607 | -102.0331
22.5644 | -102.0037
22.5681 | -101.9754
22.5718 | -101.9473
22.5755 | -101.919
22.5793 | -101.8935
22.583 | -101.8698
22.5867 | -101.847
22.5904 | -101.8239
22.5941 | -101.8004
22.5979 | -101.7783
22.6016 | -101.7561
22.6053 | -101.7367
22.609 | -101.719
22.6127 | -101.7036
22.6165 | -101.69
22.6202 | -101.6772
22.6239 | -101.6652
22.6276 | -101.6562
22.6313 | -101.6488
22.6351 | -101.6431
22.6388 | -101.6388
22.6425 | -101.6372
22.6462 | -101.6371
22.6499 | -101.6376
22.6536 | -101.6383
22.6574 | -101.6399
22.6611 | -101.6443
22.6648 | -101.6509
22.6685 | -101.6602
22.6722 | -101.6701
22.676 | -101.6814
22.6797 | -101.6947
22.6834 | -101.7073
22.6871 | -101.7205
22.6908 | -101.7341
22.6946 | -101.7494
22.6983 | -101.7684
22.702 | -101.7883
22.7057 | -101.8071
22.7094 | -101.8272
22.7132 | -101.8468
22.7169 | -101.8645
22.7206 | -101.8816
22.7243 | -101.8995
22.728 | -101.917
22.7318 | -101.9342
22.7355 | -101.9517
22.7392 | -101.9691
22.7429 | -101.9834
22.7466 | -101.9979
22.7503 | -102.0116
22.7541 | -102.0254
22.7578 | -102.0376
22.7615 | -102.0478
22.7652 | -102.0582
22.7689 | -102.0707
22.7727 | -102.0849
22.7764 | -102.0975
22.7801 | -102.1075
22.7838 | -102.1167
22.7875 | -102.1265
22.7913 | -102.1342
22.795 | -102.1421
22.7987 | -102.149
22.8024 | -102.1536
22.8061 | -102.1596
22.8099 | -102.1683
22.8136 | -102.1753
22.8173 | -102.1793
22.821 | -102.1841
22.8247 | -102.189
22.8285 | -102.1939
22.8322 | -102.199
22.8359 | -102.2054
22.8396 | -102.2099
22.8433 | -102.2154
22.847 | -102.2199
22.8508 | -102.225
22.8545 | -102.2319
22.8582 | -102.2349
22.8619 | -102.2399
22.8656 | -102.2468
22.8694 | -102.2546
22.8731 | -102.2612
22.8768 | -102.2662
22.8805 | -102.2734
22.8842 | -102.2819
22.888 | -102.2907
22.8917 | -102.2969
22.8954 | -102.3048
22.8991 | -102.3152
22.9028 | -102.3248
22.9066 | -102.3317
22.9103 | -102.3385
22.914 | -102.3464
22.9177 | -102.3536
22.9214 | -102.3621
22.9252 | -102.3732
22.9289 | -102.3821
22.9326 | -102.3909
22.9363 | -102.3997
22.94 | -102.4088
22.9437 | -102.4168
22.9475 | -102.4256
22.9512 | -102.4341
22.9549 | -102.4433
22.9586 | -102.4527
22.9623 | -102.4617
22.9661 | -102.4687
22.9698 | -102.472
22.9735 | -102.4776
22.9772 | -102.4846
22.9809 | -102.4895
22.9847 | -102.4933
22.9884 | -102.4953
22.9921 | -102.4952
22.9958 | -102.4973
22.9995 | -102.4991
23.0033 | -102.4997
23.007 | -102.5016
23.0107 | -102.5021
23.0144 | -102.502
23.0181 | -102.5004
23.0219 | -102.4945
23.0256 | -102.4888
23.0293 | -102.4837
23.033 | -102.4762
23.0367 | -102.4687
23.0404 | -102.4628
23.0442 | -102.4558
23.0479 | -102.4477
23.0516 | -102.4394
23.0553 | -102.4296
23.059 | -102.4189
23.0628 | -102.4083
23.0665 | -102.3975
23.0702 | -102.3841
23.0739 | -102.3722
23.0776 | -102.3592
23.0814 | -102.346
23.0851 | -102.3346
23.0888 | -102.3238
23.0925 | -102.3129
23.0962 | -102.2994
23.1 | -102.2855
23.1037 | -102.2688
23.1074 | -102.2542
23.1111 | -102.2433
23.1148 | -102.2312
23.1185 | -102.2199
23.1223 | -102.208
23.126 | -102.1979
23.1297 | -102.1882
23.1334 | -102.1797
23.1371 | -102.1712
23.1409 | -102.1633
23.1446 | -102.1553
23.1483 | -102.1482
23.152 | -102.1404
23.1557 | -102.1334
23.1595 | -102.1277
23.1632 | -102.1218
23.1669 | -102.1183
23.1706 | -102.1156
23.1743 | -102.1143
23.1781 | -102.1137
23.1818 | -102.1132
23.1855 | -102.1125
23.1892 | -102.1118
23.1929 | -102.1098
23.1967 | -102.1076
23.2004 | -102.1077
23.2041 | -102.11
23.2078 | -102.1122
23.2115 | -102.1135
23.2152 | -102.1143
23.219 | -102.1186
23.2227 | -102.1231
23.2264 | -102.1275
23.2301 | -102.1299
23.2338 | -102.1344
23.2376 | -102.1403
23.2413 | -102.1473
23.245 | -102.1553
23.2487 | -102.1636
23.2524 | -102.1719
23.2562 | -102.1813
23.2599 | -102.1904
23.2636 | -102.1993
23.2673 | -102.2073
23.271 | -102.2171
23.2748 | -102.2272
23.2785 | -102.2376
23.2822 | -102.2472
23.2859 | -102.2547
23.2896 | -102.2617
23.2934 | -102.2694
23.2971 | -102.2833
23.3008 | -102.2962
23.3045 | -102.3066
23.3082 | -102.3174
23.3119 | -102.3269
23.3157 | -102.3389
23.3194 | -102.3504
23.3231 | -102.3594
23.3268 | -102.3687
23.3305 | -102.3768
23.3343 | -102.3846
23.338 | -102.3944
23.3417 | -102.4033
23.3454 | -102.4124
23.3491 | -102.4229
23.3529 | -102.4325
23.3566 | -102.4402
23.3603 | -102.4499
23.364 | -102.46
23.3677 | -102.4675
23.3715 | -102.4736
23.3752 | -102.4798
23.3789 | -102.4851
23.3826 | -102.4909
23.3863 | -102.4972
23.3901 | -102.5022
23.3938 | -102.507
23.3975 | -102.5129
23.4012 | -102.5189
23.4049 | -102.5248
23.4086 | -102.5315
23.4124 | -102.5372
23.4161 | -102.5425
23.4198 | -102.5462
23.4235 | -102.5501
23.4272 | -102.5562
23.431 | -102.5622
23.4347 | -102.5698
23.4384 | -102.575
23.4421 | -102.5781
23.4458 | -102.5816
23.4496 | -102.5862
23.4533 | -102.5927
23.457 | -102.5974
23.4607 | -102.6033
23.4644 | -102.6079
23.4682 | -102.6124
23.4719 | -102.6174
23.4756 | -102.6234
23.4793 | -102.6298
23.483 | -102.6397
23.4868 | -102.6502
23.4905 | -102.6596
23.4942 | -102.6692
23.4979 | -102.6786
23.5016 | -102.6887
23.5053 | -102.6987
23.5091 | -102.7095
23.5128 | -102.7214
23.5165 | -102.7335
23.5202 | -102.7455
23.5239 | -102.7605
23.5277 | -102.7791
23.5314 | -102.7975
23.5351 | -102.8123
23.5388 | -102.8267
23.5425 | -102.8441
23.5463 | -102.8631
23.55 | -102.8797
23.5537 | -102.8959
23.5574 | -102.9125
23.5611 | -102.9337
23.5649 | -102.954
23.5686 | -102.9727
23.5723 | -102.9898
23.576 | -103.0076
23.5797 | -103.0259
23.5834 | -103.0444
23.5872 | -103.0637
23.5909 | -103.0827
23.5946 | -103.1001
23.5983 | -103.1214
23.602 | -103.1418
23.6058 | -103.1562
23.6095 | -103.1727
23.6132 | -103.19
23.6169 | -103.2069
23.6206 | -103.2253
23.6244 | -103.2428
23.6281 | -103.2597
23.6318 | -103.2716
23.6355 | -103.2834
23.6392 | -103.2941
23.643 | -103.3042
23.6467 | -103.3154
23.6504 | -103.324
23.6541 | -103.3337
23.6578 | -103.3426
23.6616 | -103.3487
23.6653 | -103.3562
23.669 | -103.3639
23.6727 | -103.3694
23.6764 | -103.3757
23.6801 | -103.3812
23.6839 | -103.3848
23.6876 | -103.3875
23.6913 | -103.3923
23.695 | -103.3966
23.6987 | -103.399
23.7025 | -103.4015
23.7062 | -103.4049
23.7099 | -103.4058
23.7136 | -103.4093
23.7173 | -103.4109
23.7211 | -103.4105
23.7248 | -103.4128
23.7285 | -103.4146
23.7322 | -103.4152
23.7359 | -103.4166
23.7397 | -103.4188
23.7434 | -103.4195
23.7471 | -103.42
23.7508 | -103.4195
23.7545 | -103.4205
23.7583 | -103.4198
23.762 | -103.419
23.7657 | -103.4186
23.7694 | -103.4205
23.7731 | -103.425
23.7768 | -103.4289
23.7806 | -103.432
23.7843 | -103.435
23.788 | -103.4366
23.7917 | -103.44
23.7954 | -103.4433
23.7992 | -103.446
23.8029 | -103.4486
23.8066 | -103.4507
23.8103 | -103.4529
23.814 | -103.4547
23.8178 | -103.4559
23.8215 | -103.4581
23.8252 | -103.4593
23.8289 | -103.4622
23.8326 | -103.4665
23.8364 | -103.4699
23.8401 | -103.4729
23.8438 | -103.4745
23.8475 | -103.4738
23.8512 | -103.4754
23.855 | -103.4771
23.8587 | -103.4777
23.8624 | -103.4745
23.8661 | -103.4684
23.8698 | -103.4659
23.8735 | -103.464
23.8773 | -103.4599
23.881 | -103.4552
23.8847 | -103.4523
23.8884 | -103.4484
23.8921 | -103.4449
23.8959 | -103.4403
23.8996 | -103.4362
23.9033 | -103.4313
23.907 | -103.4268
23.9107 | -103.4245
23.9145 | -103.4209
23.9182 | -103.4155
23.9219 | -103.4093
23.9256 | -103.4035
23.9293 | -103.3998
23.9331 | -103.394
23.9368 | -103.3873
23.9405 | -103.382
23.9442 | -103.3757
23.9479 | -103.3709
23.9517 | -103.3639
23.9554 | -103.3578
23.9591 | -103.354
23.9628 | -103.3517
23.9665 | -103.3481
23.9702 | -103.3444
23.974 | -103.3402
23.9777 | -103.3368
23.9814 | -103.3331
23.9851 | -103.3276
23.9888 | -103.323
23.9926 | -103.3188
23.9963 | -103.3152
24 | -103.3138
24.0037 | -103.3115
24.0074 | -103.3083
24.0112 | -103.3068
24.0149 | -103.3065
24.0186 | -103.3062
24.0223 | -103.307
24.026 | -103.3097
24.0298 | -103.3117
24.0335 | -103.3126
24.0372 | -103.3147
24.0409 | -103.3177
24.0446 | -103.3213
24.0483 | -103.3267
24.0521 | -103.3296
24.0558 | -103.3348
24.0595 | -103.3407
24.0632 | -103.3441
24.0669 | -103.3468
24.0707 | -103.3491
24.0744 | -103.3523
24.0781 | -103.3577
24.0818 | -103.3643
24.0855 | -103.3711
24.0893 | -103.377
24.093 | -103.3837
24.0967 | -103.3921
24.1004 | -103.3979
24.1041 | -103.4041
24.1079 | -103.4084
24.1116 | -103.4125
24.1153 | -103.4171
24.119 | -103.4226
24.1227 | -103.4294
24.1265 | -103.4366
24.1302 | -103.4426
24.1339 | -103.4469
24.1376 | -103.4518
24.1413 | -103.4586
24.145 | -103.464
24.1488 | -103.4702
24.1525 | -103.4762
24.1562 | -103.4815
24.1599 | -103.4851
24.1636 | -103.4893
24.1674 | -103.4955
24.1711 | -103.5008
24.1748 | -103.5054
24.1785 | -103.5092
24.1822 | -103.5141
24.186 | -103.5189
24.1897 | -103.5255
24.1934 | -103.5303
24.1971 | -103.5342
24.2008 | -103.5372
24.2046 | -103.5426
24.2083 | -103.5505
24.212 | -103.5552
24.2157 | -103.559
24.2194 | -103.5624
24.2232 | -103.5652
24.2269 | -103.5686
24.2306 | -103.5723
24.2343 | -103.5758
24.238 | -103.578
24.2417 | -103.5787
24.2455 | -103.5813
24.2492 | -103.5825
24.2529 | -103.5801
24.2566 | -103.5786
24.2603 | -103.5769
24.2641 | -103.5762
24.2678 | -103.5745
24.2715 | -103.5718
24.2752 | -103.5648
24.2789 | -103.5578
24.2827 | -103.5524
24.2864 | -103.5458
24.2901 | -103.5358
24.2938 | -103.5288
24.2975 | -103.5231
24.3013 | -103.516
24.305 | -103.5065
24.3087 | -103.4942
24.3124 | -103.4828
24.3161 | -103.4728
24.3199 | -103.4625
24.3236 | -103.451
24.3273 | -103.4385
24.331 | -103.4251
24.3347 | -103.4094
24.3384 | -103.3931
24.3422 | -103.3773
24.3459 | -103.3626
24.3496 | -103.345
24.3533 | -103.3268
24.357 | -103.3083
24.3608 | -103.2897
24.3645 | -103.2706
24.3682 | -103.2508
24.3719 | -103.2289
24.3756 | -103.208
24.3794 | -103.1874
24.3831 | -103.1694
24.3868 | -103.1528
24.3905 | -103.136
24.3942 | -103.1188
24.398 | -103.0999
24.4017 | -103.082
24.4054 | -103.0659
24.4091 | -103.0506
24.4128 | -103.0357
24.4166 | -103.0225
24.4203 | -103.0091
24.424 | -102.9957
24.4277 | -102.9831
24.4314 | -102.9714
24.4351 | -102.9594
24.4389 | -102.9477
24.4426 | -102.9399
24.4463 | -102.9322
24.45 | -102.9243
24.4537 | -102.9168
24.4575 | -102.91
24.4612 | -102.9051
24.4649 | -102.9005
24.4686 | -102.897
24.4723 | -102.8954
24.4761 | -102.8951
24.4798 | -102.8967
24.4835 | -102.8984
24.4872 | -102.9003
24.4909 | -102.9035
24.4947 | -102.9073
24.4984 | -102.9115
24.5021 | -102.915
24.5058 | -102.9204
24.5095 | -102.9263
24.5132 | -102.933
24.517 | -102.9399
24.5207 | -102.9449
24.5244 | -102.9512
24.5281 | -102.9592
24.5318 | -102.968
24.5356 | -102.9755
24.5393 | -102.9846
24.543 | -102.9938
24.5467 | -103.0038
24.5504 | -103.0137
24.5542 | -103.0224
24.5579 | -103.0331
24.5616 | -103.0445
24.5653 | -103.053
24.569 | -103.0567
24.5728 | -103.0597
24.5765 | -103.0642
24.5802 | -103.067
24.5839 | -103.0643
24.5876 | -103.0621
24.5914 | -103.0522
24.5951 | -103.0271
24.5988 | -102.9942
24.6025 | -102.9698
24.6062 | -102.9421
24.6099 | -102.9052
24.6137 | -102.8721
24.6174 | -102.8417
24.6211 | -102.8081
24.6248 | -102.7703
24.6285 | -102.7277
24.6323 | -102.6838
24.636 | -102.6383
24.6397 | -102.5876
24.6434 | -102.5333
24.6471 | -102.4757
24.6509 | -102.4159
24.6546 | -102.3557
24.6583 | -102.2919
24.662 | -102.2238
24.6657 | -102.1541
24.6695 | -102.0822
24.6732 | -102.0094
24.6769 | -101.9344
24.6806 | -101.8428
24.6843 | -101.7699
24.6881 | -101.6977
24.6918 | -101.6247
24.6955 | -101.5509
24.6992 | -101.4789
24.7029 | -101.4074
24.7066 | -101.3375
24.7104 | -101.2705
24.7141 | -101.205
24.7178 | -101.1424
24.7215 | -101.0835
24.7252 | -101.0272
24.729 | -100.9735
24.7327 | -100.9226
24.7364 | -100.8755
24.7401 | -100.8332
24.7438 | -100.7943
24.7476 | -100.7597
24.7513 | -100.7297
24.755 | -100.7045
24.7587 | -100.6839
24.7624 | -100.6675
24.7662 | -100.655
24.7699 | -100.647
24.7736 | -100.6453
24.7773 | -100.6489
24.781 | -100.6574
24.7848 | -100.6705
24.7885 | -100.6878
24.7922 | -100.7086
24.7959 | -100.7342
24.7996 | -100.7659
24.8033 | -100.8019
24.8071 | -100.8417
24.8108 | -100.8846
24.8145 | -100.9316
24.8182 | -100.9821
24.8219 | -101.0359
24.8257 | -101.0929
24.8294 | -101.1525
24.8331 | -101.2149
24.8368 | -101.2803
24.8405 | -101.3487
24.8443 | -101.4191
24.848 | -101.4905
24.8517 | -101.5647
24.8554 | -101.6387
24.8591 | -101.713
24.8629 | -101.7886
24.8666 | -101.8644
24.8703 | -101.9406
24.874 | -102.0163
24.8777 | -102.0898
24.8815 | -102.1621
24.8852 | -102.233
24.8889 | -102.3028
24.8926 | -102.3718
24.8963 | -102.4378
24.9 | -102.5012
24.9038 | -102.5636
24.9075 | -102.6221
24.9112 | -102.6764
24.9149 | -102.7313
24.9186 | -102.789
24.9224 | -102.8419
24.9261 | -102.889
24.9298 | -102.9327
24.9335 | -102.9783
24.9372 | -103.0179
24.941 | -103.0603
24.9447 | -103.1134
24.9484 | -103.1656
24.9521 | -103.1992
24.9558 | -103.2289
24.9596 | -103.2603
24.9633 | -103.2805
24.967 | -103.2924
24.9707 | -103.3055
24.9744 | -103.3161
24.9781 | -103.3248
24.9819 | -103.3298
24.9856 | -103.3309
24.9893 | -103.3326
24.993 | -103.3349
24.9967 | -103.3367
25.0005 | -103.3383
25.0042 | -103.3396
25.0079 | -103.3415
25.0116 | -103.3434
25.0153 | -103.3448
25.0191 | -103.348
25.0228 | -103.35
25.0265 | -103.3518
25.0302 | -103.3536
25.0339 | -103.3544
25.0377 | -103.3558
25.0414 | -103.3559
25.0451 | -103.3568
25.0488 | -103.3579
25.0525 | -103.3591
25.0563 | -103.3595
25.06 | -103.3592
25.0637 | -103.3576
25.0674 | -103.3582
25.0711 | -103.3583
25.0748 | -103.3582
25.0786 | -103.3579
25.0823 | -103.3575
25.086 | -103.3569
25.0897 | -103.3544
25.0934 | -103.3511
25.0972 | -103.3508
25.1009 | -103.3515
25.1046 | -103.3519
25.1083 | -103.3507
25.112 | -103.35
25.1158 | -103.3513
25.1195 | -103.3523
25.1232 | -103.3538
25.1269 | -103.3536
25.1306 | -103.3535
25.1344 | -103.355
25.1381 | -103.3556
25.1418 | -103.3572
25.1455 | -103.3601
25.1492 | -103.3633
25.153 | -103.365
25.1567 | -103.3658
25.1604 | -103.3676
25.1641 | -103.3698
25.1678 | -103.3721
25.1715 | -103.375
25.1753 | -103.3776
25.179 | -103.3808
25.1827 | -103.3861
25.1864 | -103.3925
25.1901 | -103.3988
25.1939 | -103.4034
25.1976 | -103.4087
25.2013 | -103.4142
25.205 | -103.4187
25.2087 | -103.4256
25.2125 | -103.4339
25.2162 | -103.4425
25.2199 | -103.449
25.2236 | -103.4566
25.2273 | -103.465
25.2311 | -103.4736
25.2348 | -103.4833
25.2385 | -103.4936
25.2422 | -103.5034
25.2459 | -103.5105
25.2497 | -103.5162
25.2534 | -103.5252
25.2571 | -103.5362
25.2608 | -103.546
25.2645 | -103.5572
25.2682 | -103.5693
25.272 | -103.5792
25.2757 | -103.5876
25.2794 | -103.5959
25.2831 | -103.606
25.2868 | -103.6161
25.2906 | -103.6251
25.2943 | -103.6363
25.298 | -103.6443
25.3017 | -103.6516
25.3054 | -103.6587
25.3092 | -103.6654
25.3129 | -103.6718
25.3166 | -103.6774
25.3203 | -103.6839
25.324 | -103.689
25.3278 | -103.6926
25.3315 | -103.6946
25.3352 | -103.6963
25.3389 | -103.6996
25.3426 | -103.7057
25.3464 | -103.7095
25.3501 | -103.71
25.3538 | -103.7098
25.3575 | -103.7098
25.3612 | -103.7076
25.3649 | -103.7063
25.3687 | -103.7048
25.3724 | -103.7032
25.3761 | -103.7011
25.3798 | -103.6995
25.3835 | -103.6981
25.3873 | -103.6953
25.391 | -103.6919
25.3947 | -103.6877
25.3984 | -103.6834
25.4021 | -103.6792
25.4059 | -103.6738
25.4096 | -103.6696
25.4133 | -103.6655
25.417 | -103.6618
25.4207 | -103.6579
25.4245 | -103.6508
25.4282 | -103.6429
25.4319 | -103.6344
25.4356 | -103.628
25.4393 | -103.6213
25.443 | -103.6155
25.4468 | -103.6097
25.4505 | -103.6035
25.4542 | -103.5977
25.4579 | -103.5918
25.4616 | -103.5835
25.4654 | -103.5774
25.4691 | -103.572
25.4728 | -103.5655
25.4765 | -103.5609
25.4802 | -103.5564
25.484 | -103.5528
25.4877 | -103.5507
25.4914 | -103.5483
25.4951 | -103.5475
25.4988 | -103.5457
25.5026 | -103.5448
25.5063 | -103.5451
25.51 | -103.5466
25.5137 | -103.549
25.5174 | -103.5515
25.5212 | -103.5556
25.5249 | -103.5612
25.5286 | -103.5676
25.5323 | -103.5739
25.536 | -103.5802
25.5397 | -103.5832
25.5435 | -103.5861
25.5472 | -103.5919
25.5509 | -103.5986
25.5546 | -103.6048
25.5583 | -103.6131
25.5621 | -103.6218
25.5658 | -103.6302
25.5695 | -103.638
25.5732 | -103.6463
25.5769 | -103.6545
25.5807 | -103.6633
25.5844 | -103.6738
25.5881 | -103.6846
25.5918 | -103.6939
25.5955 | -103.7021
25.5993 | -103.7107
25.603 | -103.7212
25.6067 | -103.7339
25.6104 | -103.7471
25.6141 | -103.7576
25.6179 | -103.7659
25.6216 | -103.7742
25.6253 | -103.7829
25.629 | -103.7925
25.6327 | -103.8039
25.6364 | -103.8147
25.6402 | -103.8253
25.6439 | -103.8351
25.6476 | -103.846
25.6513 | -103.8584
25.655 | -103.8683
25.6588 | -103.8779
25.6625 | -103.8875
25.6662 | -103.8961
25.6699 | -103.9056
25.6736 | -103.9171
25.6774 | -103.9275
25.6811 | -103.9352
25.6848 | -103.9437
25.6885 | -103.9537
25.6922 | -103.9642
25.696 | -103.9728
25.6997 | -103.9815
25.7034 | -103.9877
25.7071 | -103.9931
25.7108 | -103.9977
25.7146 | -104.0047
25.7183 | -104.0114
25.722 | -104.0199
25.7257 | -104.0271
25.7294 | -104.0334
25.7331 | -104.0389
25.7369 | -104.0417
25.7406 | -104.0455
25.7443 | -104.0517
25.748 | -104.0579
25.7517 | -104.0624
25.7555 | -104.0653
25.7592 | -104.0688
25.7629 | -104.071
25.7666 | -104.074
25.7703 | -104.0797
25.7741 | -104.0823
25.7778 | -104.0845
25.7815 | -104.0877
25.7852 | -104.0912
25.7889 | -104.0954
25.7927 | -104.1009
25.7964 | -104.107
25.8001 | -104.1122
25.8038 | -104.1165
25.8075 | -104.1185
25.8113 | -104.1226
25.815 | -104.1271
25.8187 | -104.1298
25.8224 | -104.1314
25.8261 | -104.1354
25.8298 | -104.1376
25.8336 | -104.1379
25.8373 | -104.1392
25.841 | -104.1389
25.8447 | -104.1382
25.8484 | -104.1369
25.8522 | -104.1331
25.8559 | -104.1299
25.8596 | -104.1261
25.8633 | -104.1239
25.867 | -104.1198
25.8708 | -104.1139
25.8745 | -104.1094
25.8782 | -104.1054
25.8819 | -104.1012
25.8856 | -104.0939
25.8894 | -104.0864
25.8931 | -104.0804
25.8968 | -104.0747
25.9005 | -104.0679
25.9042 | -104.0629
25.9079 | -104.0563
25.9117 | -104.047
25.9154 | -104.0396
25.9191 | -104.0336
25.9228 | -104.0263
25.9265 | -104.0193
25.9303 | -104.0113
25.934 | -104.0031
25.9377 | -103.996
25.9414 | -103.9888
25.9451 | -103.9796
25.9489 | -103.9704
25.9526 | -103.961
25.9563 | -103.9542
25.96 | -103.9497
25.9637 | -103.9439
25.9675 | -103.9374
25.9712 | -103.93
25.9749 | -103.9225
25.9786 | -103.9148
25.9823 | -103.9094
25.9861 | -103.9062
25.9898 | -103.9036
25.9935 | -103.9012
25.9972 | -103.8995
26.0009 | -103.8979
26.0046 | -103.897
26.0084 | -103.8964
26.0121 | -103.8957
26.0158 | -103.8942
26.0195 | -103.8947
26.0232 | -103.8975
26.027 | -103.8995
26.0307 | -103.9025
26.0344 | -103.9059
26.0381 | -103.9075
26.0418 | -103.9111
26.0456 | -103.9155
26.0493 | -103.9194
26.053 | -103.9235
26.0567 | -103.9288
26.0604 | -103.9346
26.0642 | -103.938
26.0679 | -103.9428
26.0716 | -103.9483
26.0753 | -103.9548
26.079 | -103.9594
26.0828 | -103.9638
26.0865 | -103.97
26.0902 | -103.9757
26.0939 | -103.9813
26.0976 | -103.9888
26.1013 | -103.9965
26.1051 | -104.0027
26.1088 | -104.0078
26.1125 | -104.0133
26.1162 | -104.0183
26.1199 | -104.0261
26.1237 | -104.0336
26.1274 | -104.0406
26.1311 | -104.0473
26.1348 | -104.0509
26.1385 | -104.0549
26.1423 | -104.0605
26.146 | -104.0632
26.1497 | -104.0648
26.1534 | -104.0676
26.1571 | -104.0716
26.1609 | -104.0746
26.1646 | -104.0776
26.1683 | -104.0801
26.172 | -104.0828
26.1757 | -104.0877
26.1795 | -104.0902
26.1832 | -104.093
26.1869 | -104.0968
26.1906 | -104.1018
26.1943 | -104.1054
26.198 | -104.1087
26.2018 | -104.1124
26.2055 | -104.1149
26.2092 | -104.1169
26.2129 | -104.1197
26.2166 | -104.1223
26.2204 | -104.1243
26.2241 | -104.127
26.2278 | -104.1301
26.2315 | -104.1319
26.2352 | -104.1342
26.239 | -104.137
26.2427 | -104.1377
26.2464 | -104.1368
26.2501 | -104.1369
26.2538 | -104.1399
26.2576 | -104.1421
26.2613 | -104.142
26.265 | -104.1431
26.2687 | -104.1449
26.2724 | -104.1472
26.2762 | -104.1513
26.2799 | -104.155
26.2836 | -104.1571
26.2873 | -104.159
26.291 | -104.1619
26.2947 | -104.1647
26.2985 | -104.1673
26.3022 | -104.1696
26.3059 | -104.1716
26.3096 | -104.1741
26.3133 | -104.1782
26.3171 | -104.1834
26.3208 | -104.1874
26.3245 | -104.1927
26.3282 | -104.1979
26.3319 | -104.2006
26.3357 | -104.2038
26.3394 | -104.2076
26.3431 | -104.2115
26.3468 | -104.2177
26.3505 | -104.2231
26.3543 | -104.2268
26.358 | -104.2289
26.3617 | -104.2312
26.3654 | -104.2325
26.3691 | -104.232
26.3728 | -104.2314
26.3766 | -104.2337
26.3803 | -104.2382
26.384 | -104.2437
26.3877 | -104.2477
26.3914 | -104.2508
26.3952 | -104.2538
26.3989 | -104.2577
26.4026 | -104.2611
26.4063 | -104.2638
26.41 | -104.2666
26.4138 | -104.2685
26.4175 | -104.2702
26.4212 | -104.2732
26.4249 | -104.2743
26.4286 | -104.2752
26.4324 | -104.2777
26.4361 | -104.28
26.4398 | -104.2811
26.4435 | -104.2816
26.4472 | -104.2837
26.451 | -104.2857
26.4547 | -104.2866
26.4584 | -104.2859
26.4621 | -104.2847
26.4658 | -104.2846
26.4695 | -104.2786
26.4733 | -104.2728
26.477 | -104.269
26.4807 | -104.2658
26.4844 | -104.2628
26.4881 | -104.2594
26.4919 | -104.2541
26.4956 | -104.2497
26.4993 | -104.2448
26.503 | -104.2387
26.5067 | -104.2321
26.5105 | -104.2252
26.5142 | -104.2186
26.5179 | -104.2138
26.5216 | -104.2075
26.5253 | -104.1998
26.5291 | -104.1915
26.5328 | -104.1831
26.5365 | -104.1747
26.5402 | -104.1654
26.5439 | -104.1553
26.5477 | -104.1445
26.5514 | -104.1342
26.5551 | -104.123
26.5588 | -104.1114
26.5625 | -104.1009
26.5662 | -104.0905
26.57 | -104.0799
26.5737 | -104.0702
26.5774 | -104.0609
26.5811 | -104.0511
26.5848 | -104.0409
26.5886 | -104.0319
26.5923 | -104.0232
26.596 | -104.0125
26.5997 | -104.0018
26.6034 | -103.9907
26.6072 | -103.9785
26.6109 | -103.9674
26.6146 | -103.9574
26.6183 | -103.9473
26.622 | -103.9369
26.6258 | -103.9261
26.6295 | -103.9147
26.6332 | -103.9053
26.6369 | -103.8962
26.6406 | -103.8883
26.6444 | -103.8807
26.6481 | -103.8733
26.6518 | -103.8652
26.6555 | -103.857
26.6592 | -103.8488
26.6629 | -103.843
26.6667 | -103.8378
26.6704 | -103.8329
26.6741 | -103.828
26.6778 | -103.8235
26.6815 | -103.8206
26.6853 | -103.8185
26.689 | -103.8139
26.6927 | -103.8092
26.6964 | -103.8063
26.7001 | -103.8033
26.7039 | -103.7996
26.7076 | -103.7986
26.7113 | -103.7979
26.715 | -103.7972
26.7187 | -103.798
26.7225 | -103.797
26.7262 | -103.7984
26.7299 | -103.7983
26.7336 | -103.7987
26.7373 | -103.7993
26.7411 | -103.8013
26.7448 | -103.804
26.7485 | -103.8069
26.7522 | -103.8083
26.7559 | -103.807
26.7596 | -103.8055
26.7634 | -103.806
26.7671 | -103.8044
26.7708 | -103.8017
26.7745 | -103.8002
26.7782 | -103.7992
26.782 | -103.7988
26.7857 | -103.7984
26.7894 | -103.7969
26.7931 | -103.7946
26.7968 | -103.7912
26.8006 | -103.7884
26.8043 | -103.7862
26.808 | -103.7844
26.8117 | -103.783
26.8154 | -103.7818
26.8192 | -103.7794
26.8229 | -103.7756
26.8266 | -103.7735
26.8303 | -103.7704
26.834 | -103.7681
26.8377 | -103.766
26.8415 | -103.7636
26.8452 | -103.7609
26.8489 | -103.7612
26.8526 | -103.7598
26.8563 | -103.7559
26.8601 | -103.7514
26.8638 | -103.7476
26.8675 | -103.745
26.8712 | -103.7417
26.8749 | -103.7386
26.8787 | -103.7347
26.8824 | -103.7309
26.8861 | -103.7275
26.8898 | -103.7235
26.8935 | -103.7201
26.8973 | -103.7165
26.901 | -103.7138
26.9047 | -103.7117
26.9084 | -103.7082
26.9121 | -103.7036
26.9159 | -103.7012
26.9196 | -103.7005
26.9233 | -103.7003
26.927 | -103.7004
26.9307 | -103.7019
26.9344 | -103.7027
26.9382 | -103.7024
26.9419 | -103.7017
26.9456 | -103.7013
26.9493 | -103.7012
26.953 | -103.7012
26.9568 | -103.7033
26.9605 | -103.7067
26.9642 | -103.7104
26.9679 | -103.7136
26.9716 | -103.7167
26.9754 | -103.7206
26.9791 | -103.7239
26.9828 | -103.7287
26.9865 | -103.7346
26.9902 | -103.7389
26.994 | -103.7436
26.9977 | -103.7493
27.0014 | -103.7545
27.0051 | -103.7598
27.0088 | -103.7664
27.0126 | -103.772
27.0163 | -103.7769
27.02 | -103.7805
27.0237 | -103.7835
27.0274 | -103.7872
27.0311 | -103.7916
27.0349 | -103.7941
27.0386 | -103.796
27.0423 | -103.7986
27.046 | -103.7993
27.0497 | -103.7995
27.0535 | -103.8012
27.0572 | -103.8028
27.0609 | -103.8033
27.0646 | -103.8023
27.0683 | -103.8019
27.0721 | -103.8008
27.0758 | -103.7995
27.0795 | -103.799
27.0832 | -103.7981
27.0869 | -103.7946
27.0907 | -103.7904
27.0944 | -103.7858
27.0981 | -103.7791
27.1018 | -103.7738
27.1055 | -103.7672
27.1093 | -103.7613
27.113 | -103.7561
27.1167 | -103.7485
27.1204 | -103.7406
27.1241 | -103.733
27.1278 | -103.7238
27.1316 | -103.7118
27.1353 | -103.6984
27.139 | -103.6847
27.1427 | -103.6716
27.1464 | -103.6619
27.1502 | -103.6517
27.1539 | -103.6401
27.1576 | -103.6284
27.1613 | -103.6151
27.165 | -103.6011
27.1688 | -103.587
27.1725 | -103.5752
27.1762 | -103.5625
27.1799 | -103.5487
27.1836 | -103.5339
27.1874 | -103.5179
27.1911 | -103.5027
27.1948 | -103.4867
27.1985 | -103.4726
27.2022 | -103.4594
27.206 | -103.4447
27.2097 | -103.4295
27.2134 | -103.4167
27.2171 | -103.4056
27.2208 | -103.3946
27.2245 | -103.3841
27.2283 | -103.3732
27.232 | -103.3634
27.2357 | -103.355
27.2394 | -103.3466
27.2431 | -103.3372
27.2469 | -103.3286
27.2506 | -103.3225
27.2543 | -103.3179
27.258 | -103.3119
27.2617 | -103.3069
27.2655 | -103.3041
27.2692 | -103.301
27.2729 | -103.2996
27.2766 | -103.2975
27.2803 | -103.2955
27.2841 | -103.2948
27.2878 | -103.2942
27.2915 | -103.2955
27.2952 | -103.2975
27.2989 | -103.2983
27.3026 | -103.2992
27.3064 | -103.3005
27.3101 | -103.3016
27.3138 | -103.3017
27.3175 | -103.3046
27.3212 | -103.3088
27.325 | -103.312
27.3287 | -103.3169
27.3324 | -103.3233
27.3361 | -103.3306
27.3398 | -103.3365
27.3436 | -103.3413
27.3473 | -103.3458
27.351 | -103.3502
27.3547 | -103.3555
27.3584 | -103.362
27.3622 | -103.3682
27.3659 | -103.3748
27.3696 | -103.3812
27.3733 | -103.3869
27.377 | -103.3919
27.3808 | -103.3958
27.3845 | -103.4005
27.3882 | -103.4064
27.3919 | -103.4115
27.3956 | -103.4163
27.3993 | -103.4213
27.4031 | -103.4258
27.4068 | -103.4304
27.4105 | -103.4354
27.4142 | -103.4384
27.4179 | -103.4437
27.4217 | -103.4506
27.4254 | -103.4566
27.4291 | -103.4612
27.4328 | -103.4667
27.4365 | -103.4724
27.4403 | -103.4767
27.444 | -103.4787
27.4477 | -103.4813
27.4514 | -103.4852
27.4551 | -103.4896
27.4589 | -103.4933
27.4626 | -103.4956
27.4663 | -103.4984
27.47 | -103.5017
27.4737 | -103.5061
27.4775 | -103.5111
27.4812 | -103.5149
27.4849 | -103.5179
27.4886 | -103.5184
27.4923 | -103.5194
27.496 | -103.5215
27.4998 | -103.5233
27.5035 | -103.5256
27.5072 | -103.5256
27.5109 | -103.5257
27.5146 | -103.5281
27.5184 | -103.5324
27.5221 | -103.5375
27.5258 | -103.5409
27.5295 | -103.5429
27.5332 | -103.5449
27.537 | -103.5463
27.5407 | -103.547
27.5444 | -103.549
27.5481 | -103.5512
27.5518 | -103.5531
27.5556 | -103.5559
27.5593 | -103.5569
27.563 | -103.5588
27.5667 | -103.5597
27.5704 | -103.5606
27.5742 | -103.5599
27.5779 | -103.5598
27.5816 | -103.5609
27.5853 | -103.5617
27.589 | -103.5628
27.5927 | -103.5642
27.5965 | -103.5651
27.6002 | -103.5642
27.6039 | -103.5614
27.6076 | -103.5596
27.6113 | -103.559
27.6151 | -103.5582
27.6188 | -103.5565
27.6225 | -103.5556
27.6262 | -103.5553
27.6299 | -103.555
27.6337 | -103.5532
27.6374 | -103.55
27.6411 | -103.5461
27.6448 | -103.5434
27.6485 | -103.5414
27.6523 | -103.5394
27.656 | -103.5368
27.6597 | -103.5319
27.6634 | -103.5283
27.6671 | -103.5259
27.6709 | -103.5229
27.6746 | -103.5188
27.6783 | -103.5146
27.682 | -103.5124
27.6857 | -103.5091
27.6894 | -103.5062
27.6932 | -103.503
27.6969 | -103.501
27.7006 | -103.4992
27.7043 | -103.4957
27.708 | -103.493
27.7118 | -103.4925
27.7155 | -103.4913
27.7192 | -103.488
27.7229 | -103.4834
27.7266 | -103.4794
27.7304 | -103.4772
27.7341 | -103.4763
27.7378 | -103.4773
27.7415 | -103.4793
27.7452 | -103.48
27.749 | -103.4789
27.7527 | -103.4794
27.7564 | -103.4792
27.7601 | -103.4798
27.7638 | -103.4823
27.7675 | -103.4861
27.7713 | -103.4906
27.775 | -103.4941
27.7787 | -103.4955
27.7824 | -103.4974
27.7861 | -103.4996
27.7899 | -103.5036
27.7936 | -103.5084
27.7973 | -103.5124
27.801 | -103.5142
27.8047 | -103.517
27.8085 | -103.5195
27.8122 | -103.5214
27.8159 | -103.5236
27.8196 | -103.5239
27.8233 | -103.5252
27.8271 | -103.5273
27.8308 | -103.5297
27.8345 | -103.5324
27.8382 | -103.5347
27.8419 | -103.5363
27.8457 | -103.537
27.8494 | -103.5376
27.8531 | -103.5382
27.8568 | -103.5374
27.8605 | -103.5353
27.8642 | -103.5341
27.868 | -103.533
27.8717 | -103.5323
27.8754 | -103.5304
27.8791 | -103.5295
27.8828 | -103.5278
27.8866 | -103.5243
27.8903 | -103.5173
27.894 | -103.5109
27.8977 | -103.5038
27.9014 | -103.5004
27.9052 | -103.4972
27.9089 | -103.4925
27.9126 | -103.4866
27.9163 | -103.4804
27.92 | -103.4758
27.9238 | -103.4704
27.9275 | -103.4627
27.9312 | -103.4565
27.9349 | -103.4506
27.9386 | -103.4429
27.9424 | -103.4354
27.9461 | -103.4281
27.9498 | -103.4208
27.9535 | -103.4139
27.9572 | -103.4079
27.9609 | -103.4045
27.9647 | -103.3991
27.9684 | -103.3961
27.9721 | -103.3923
27.9758 | -103.3876
27.9795 | -103.3825
27.9833 | -103.3763
27.987 | -103.3706
27.9907 | -103.3664
27.9944 | -103.3637
27.9981 | -103.3623
28.0019 | -103.3578