## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 225 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 228
Hood | 02-92B/P | 167
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1014.2 | hPa |  | PRC PRO 01
Humidity | 40.2 | %/Hr
Internal temperature of calibration bench | 24.01 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1015 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 23.5 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.93 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.02 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3061 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 15/05/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 16/05/2018
Engineer | Anthony HUE | Approval |  | 18/05/2018
CEO | Robert GLEMOT | Approval |  | 18/05/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.347 | 57.4683
0.012209 | -36.0014 | 38.7775
0.018313 | -35.0129 | 28.6755
0.024417 | -34.5723 | 22.1503
0.030521 | -34.4082 | 18.1311
0.036626 | -34.2614 | 15.0524
0.04273 | -34.1768 | 12.9206
0.048834 | -34.122 | 11.2545
0.054938 | -34.0784 | 9.9768
0.061043 | -34.028 | 8.9306
0.067147 | -34.0121 | 8.0532
0.073251 | -34.001 | 7.2966
0.079355 | -33.9972 | 6.7479
0.08546 | -33.9981 | 6.3314
0.091564 | -33.9931 | 5.9536
0.097668 | -33.9697 | 5.6019
0.10377 | -33.9372 | 5.203
0.10988 | -33.9169 | 4.817
0.11598 | -33.9124 | 4.4645
0.12209 | -33.9259 | 4.2018
0.12819 | -33.95 | 3.9987
0.13429 | -33.9729 | 3.8075
0.1404 | -33.9878 | 3.6051
0.1465 | -34.0019 | 3.4055
0.15261 | -34.011 | 3.2101
0.15871 | -34.0121 | 3.0243
0.16482 | -34.0033 | 2.8377
0.17092 | -33.9802 | 2.7061
0.17702 | -33.9626 | 2.5651
0.18313 | -33.9458 | 2.4475
0.18923 | -33.9334 | 2.3349
0.19534 | -33.9274 | 2.244
0.20144 | -33.9273 | 2.1597
0.20754 | -33.9324 | 2.1104
0.21365 | -33.9357 | 2.0542
0.21975 | -33.9406 | 1.9997
0.22586 | -33.948 | 1.9397
0.23196 | -33.955 | 1.8631
0.23807 | -33.9608 | 1.7759
0.24417 | -33.9601 | 1.6722
0.25027 | -33.9608 | 1.5803
0.25638 | -33.9592 | 1.4824
0.26248 | -33.9563 | 1.3946
0.26859 | -33.9535 | 1.3114
0.27469 | -33.9529 | 1.2317
0.2808 | -33.9584 | 1.1813
0.2869 | -33.9627 | 1.1233
0.293 | -33.9679 | 1.0751
0.29911 | -33.974 | 1.0423
0.30521 | -33.9796 | 1.0119
0.31132 | -33.9836 | 0.9925
0.31742 | -33.9875 | 0.96675
0.32353 | -33.9911 | 0.94098
0.32963 | -33.9938 | 0.91549
0.33573 | -33.9954 | 0.88705
0.34184 | -33.9973 | 0.85748
0.34794 | -33.9992 | 0.81406
0.35405 | -34.0012 | 0.77512
0.36015 | -34.0031 | 0.73424
0.36626 | -34.0042 | 0.69217
0.37236 | -34.0053 | 0.6445
0.37846 | -34.006 | 0.59481
0.38457 | -34.0043 | 0.55317
0.39067 | -34.0043 | 0.51272
0.39678 | -34.0023 | 0.46988
0.40288 | -33.9996 | 0.43157
0.40899 | -33.996 | 0.3934
0.41509 | -33.9921 | 0.35397
0.42119 | -33.9873 | 0.31445
0.4273 | -33.9838 | 0.27541
0.4334 | -33.9807 | 0.23602
0.43951 | -33.9777 | 0.19567
0.44561 | -33.9758 | 0.15621
0.45172 | -33.9766 | 0.12066
0.45782 | -33.9761 | 0.084133
0.46392 | -33.9768 | 0.051415
0.47003 | -33.978 | 0.019681
0.47613 | -33.9798 | -0.0073193
0.48224 | -33.982 | -0.031625
0.48834 | -33.9855 | -0.045005
0.49445 | -33.9881 | -0.06215
0.50055 | -33.9914 | -0.078882
0.50665 | -33.9941 | -0.08934
0.51276 | -33.9964 | -0.10075
0.51886 | -33.9986 | -0.10996
0.52497 | -33.9989 | -0.1145
0.53107 | -34.0002 | -0.11907
0.53717 | -34.0014 | -0.12548
0.54328 | -34.0022 | -0.12952
0.54938 | -34.0026 | -0.13674
0.55549 | -34.0024 | -0.1459
0.56159 | -34.0018 | -0.16353
0.5677 | -34.0015 | -0.17878
0.5738 | -34.0005 | -0.19544
0.5799 | -34.0002 | -0.2118
0.58601 | -34.0002 | -0.23619
0.59211 | -34.0003 | -0.26982
0.59822 | -34.0003 | -0.29505
0.60432 | -34.0006 | -0.32304
0.61043 | -34.001 | -0.35401
0.61653 | -34.0017 | -0.38349
0.62263 | -34.0019 | -0.41529
0.62874 | -34.0031 | -0.44315
0.63484 | -34.004 | -0.47308
0.64095 | -34.0052 | -0.50188
0.64705 | -34.0064 | -0.52983
0.65316 | -34.0076 | -0.55526
0.65926 | -34.0089 | -0.58115
0.66536 | -34.0095 | -0.60056
0.67147 | -34.0104 | -0.62148
0.67757 | -34.0115 | -0.64289
0.68368 | -34.0121 | -0.66256
0.68978 | -34.0131 | -0.68145
0.69589 | -34.0137 | -0.69879
0.70199 | -34.0139 | -0.716
0.70809 | -34.0142 | -0.73485
0.7142 | -34.0145 | -0.75145
0.7203 | -34.0148 | -0.76909
0.72641 | -34.0147 | -0.78552
0.73251 | -34.0152 | -0.80324
0.73862 | -34.0152 | -0.82251
0.74472 | -34.0159 | -0.83878
0.75082 | -34.0159 | -0.85499
0.75693 | -34.0159 | -0.87366
0.76303 | -34.0156 | -0.89111
0.76914 | -34.0149 | -0.91426
0.77524 | -34.0145 | -0.936
0.78135 | -34.0141 | -0.95573
0.78745 | -34.0139 | -0.97714
0.79355 | -34.0136 | -0.99808
0.79966 | -34.0134 | -1.0207
0.80576 | -34.013 | -1.0475
0.81187 | -34.0123 | -1.0722
0.81797 | -34.0116 | -1.0982
0.82408 | -34.0112 | -1.1211
0.83018 | -34.0106 | -1.148
0.83628 | -34.0102 | -1.1759
0.84239 | -34.0093 | -1.1998
0.84849 | -34.0084 | -1.2229
0.8546 | -34.0079 | -1.2468
0.8607 | -34.0072 | -1.2716
0.86681 | -34.0067 | -1.2971
0.87291 | -34.0057 | -1.3199
0.87901 | -34.005 | -1.3401
0.88512 | -34.0045 | -1.3607
0.89122 | -34.004 | -1.3816
0.89733 | -34.0037 | -1.4003
0.90343 | -34.0032 | -1.4196
0.90953 | -34.0035 | -1.4349
0.91564 | -34.0031 | -1.4503
0.92174 | -34.0028 | -1.4642
0.92785 | -34.0023 | -1.4781
0.93395 | -34.0024 | -1.4906
0.94006 | -34.0024 | -1.5056
0.94616 | -34.0026 | -1.512
0.95226 | -34.0028 | -1.5249
0.95837 | -34.003 | -1.537
0.96447 | -34.0034 | -1.5485
0.97058 | -34.0037 | -1.5591
0.97668 | -34.0038 | -1.5679
0.98279 | -34.0041 | -1.5724
0.98889 | -34.0047 | -1.5798
0.99499 | -34.0049 | -1.5887
1.0011 | -34.0054 | -1.5959
1.0072 | -34.0057 | -1.6012
1.0133 | -34.006 | -1.6073
1.0194 | -34.0065 | -1.6129
1.0255 | -34.0066 | -1.6198
1.0316 | -34.0065 | -1.6252
1.0377 | -34.0065 | -1.6309
1.0438 | -34.0063 | -1.6358
1.0499 | -34.0055 | -1.6407
1.056 | -34.005 | -1.6449
1.0621 | -34.0045 | -1.65
1.0682 | -34.0045 | -1.6557
1.0743 | -34.004 | -1.6619
1.0805 | -34.0039 | -1.6679
1.0866 | -34.0031 | -1.6775
1.0927 | -34.0026 | -1.6841
1.0988 | -34.0024 | -1.6938
1.1049 | -34.0017 | -1.7006
1.111 | -34.0009 | -1.7103
1.1171 | -34.0002 | -1.7189
1.1232 | -33.9993 | -1.7318
1.1293 | -33.9987 | -1.7432
1.1354 | -33.9981 | -1.7543
1.1415 | -33.9973 | -1.7667
1.1476 | -33.9966 | -1.779
1.1537 | -33.9958 | -1.7913
1.1598 | -33.9958 | -1.8083
1.1659 | -33.9956 | -1.8203
1.172 | -33.9952 | -1.833
1.1781 | -33.9949 | -1.8464
1.1842 | -33.9946 | -1.8614
1.1903 | -33.9953 | -1.8781
1.1964 | -33.9954 | -1.8923
1.2025 | -33.9954 | -1.9067
1.2086 | -33.9958 | -1.9226
1.2147 | -33.9963 | -1.9378
1.2209 | -33.9969 | -1.9523
1.227 | -33.9985 | -1.9686
1.2331 | -33.9993 | -1.9834
1.2392 | -34.0002 | -1.9998
1.2453 | -34.0013 | -2.0129
1.2514 | -34.0025 | -2.0268
1.2575 | -34.0038 | -2.0419
1.2636 | -34.0057 | -2.0557
1.2697 | -34.0072 | -2.0695
1.2758 | -34.0085 | -2.0846
1.2819 | -34.0099 | -2.0999
1.288 | -34.0113 | -2.1147
1.2941 | -34.013 | -2.1281
1.3002 | -34.0147 | -2.1393
1.3063 | -34.0161 | -2.1517
1.3124 | -34.0177 | -2.1641
1.3185 | -34.0193 | -2.177
1.3246 | -34.0208 | -2.1902
1.3307 | -34.0225 | -2.2023
1.3368 | -34.024 | -2.2122
1.3429 | -34.0253 | -2.2234
1.349 | -34.0267 | -2.234
1.3551 | -34.028 | -2.2471
1.3613 | -34.0294 | -2.2578
1.3674 | -34.0304 | -2.2677
1.3735 | -34.0315 | -2.2775
1.3796 | -34.0326 | -2.2871
1.3857 | -34.0336 | -2.2962
1.3918 | -34.0344 | -2.3074
1.3979 | -34.0355 | -2.3179
1.404 | -34.036 | -2.3285
1.4101 | -34.0368 | -2.3387
1.4162 | -34.0375 | -2.3484
1.4223 | -34.0381 | -2.3592
1.4284 | -34.0387 | -2.3691
1.4345 | -34.0393 | -2.3799
1.4406 | -34.0397 | -2.3898
1.4467 | -34.0402 | -2.3995
1.4528 | -34.0407 | -2.4091
1.4589 | -34.0413 | -2.4198
1.465 | -34.0415 | -2.4283
1.4711 | -34.0423 | -2.4378
1.4772 | -34.0426 | -2.4489
1.4833 | -34.043 | -2.4604
1.4894 | -34.0435 | -2.4717
1.4955 | -34.044 | -2.4833
1.5016 | -34.0444 | -2.495
1.5078 | -34.0447 | -2.5082
1.5139 | -34.0452 | -2.5201
1.52 | -34.0455 | -2.5328
1.5261 | -34.0459 | -2.544
1.5322 | -34.0459 | -2.5552
1.5383 | -34.0461 | -2.5669
1.5444 | -34.0465 | -2.5797
1.5505 | -34.0467 | -2.5918
1.5566 | -34.047 | -2.6039
1.5627 | -34.0474 | -2.6158
1.5688 | -34.0479 | -2.6278
1.5749 | -34.048 | -2.6396
1.581 | -34.0483 | -2.6552
1.5871 | -34.0487 | -2.6695
1.5932 | -34.0491 | -2.6817
1.5993 | -34.0493 | -2.6968
1.6054 | -34.0495 | -2.7108
1.6115 | -34.0497 | -2.7261
1.6176 | -34.05 | -2.7419
1.6237 | -34.0502 | -2.7551
1.6298 | -34.0505 | -2.77
1.6359 | -34.0508 | -2.7839
1.642 | -34.0511 | -2.7983
1.6482 | -34.0514 | -2.8144
1.6543 | -34.0518 | -2.8298
1.6604 | -34.0522 | -2.8442
1.6665 | -34.0523 | -2.8587
1.6726 | -34.0526 | -2.8744
1.6787 | -34.0531 | -2.8909
1.6848 | -34.0531 | -2.9068
1.6909 | -34.0535 | -2.9223
1.697 | -34.0537 | -2.9381
1.7031 | -34.0539 | -2.9531
1.7092 | -34.054 | -2.9676
1.7153 | -34.0543 | -2.9836
1.7214 | -34.0541 | -2.9985
1.7275 | -34.0541 | -3.0138
1.7336 | -34.0541 | -3.0293
1.7397 | -34.0542 | -3.0441
1.7458 | -34.0544 | -3.0603
1.7519 | -34.0543 | -3.0761
1.758 | -34.0543 | -3.0906
1.7641 | -34.0544 | -3.1051
1.7702 | -34.0543 | -3.1199
1.7763 | -34.0545 | -3.1337
1.7824 | -34.0545 | -3.1493
1.7885 | -34.0544 | -3.1626
1.7947 | -34.0546 | -3.1772
1.8008 | -34.0546 | -3.1902
1.8069 | -34.0546 | -3.2039
1.813 | -34.0547 | -3.2175
1.8191 | -34.0545 | -3.2309
1.8252 | -34.0544 | -3.2438
1.8313 | -34.0544 | -3.2576
1.8374 | -34.0544 | -3.2709
1.8435 | -34.0546 | -3.2827
1.8496 | -34.0544 | -3.296
1.8557 | -34.0543 | -3.3098
1.8618 | -34.0541 | -3.3208
1.8679 | -34.0542 | -3.3334
1.874 | -34.0542 | -3.3461
1.8801 | -34.0541 | -3.3584
1.8862 | -34.0542 | -3.3707
1.8923 | -34.0541 | -3.3823
1.8984 | -34.0538 | -3.3923
1.9045 | -34.0538 | -3.4032
1.9106 | -34.0536 | -3.4145
1.9167 | -34.0534 | -3.4251
1.9228 | -34.0534 | -3.4356
1.9289 | -34.053 | -3.444
1.9351 | -34.0531 | -3.4534
1.9412 | -34.0531 | -3.4637
1.9473 | -34.0531 | -3.4751
1.9534 | -34.0529 | -3.4861
1.9595 | -34.0527 | -3.4973
1.9656 | -34.0522 | -3.5071
1.9717 | -34.0519 | -3.5172
1.9778 | -34.0518 | -3.5286
1.9839 | -34.0517 | -3.5399
1.99 | -34.0518 | -3.5499
1.9961 | -34.0516 | -3.5595
2.0022 | -34.0514 | -3.5693
2.0083 | -34.0511 | -3.5792
2.0144 | -34.0511 | -3.5889
2.0205 | -34.051 | -3.6003
2.0266 | -34.051 | -3.6118
2.0327 | -34.051 | -3.6226
2.0388 | -34.051 | -3.6362
2.0449 | -34.051 | -3.6463
2.051 | -34.051 | -3.6562
2.0571 | -34.0507 | -3.6674
2.0632 | -34.0506 | -3.6777
2.0693 | -34.0506 | -3.6881
2.0754 | -34.0507 | -3.6997
2.0816 | -34.0508 | -3.711
2.0877 | -34.0507 | -3.7219
2.0938 | -34.0508 | -3.7334
2.0999 | -34.0507 | -3.7446
2.106 | -34.051 | -3.7562
2.1121 | -34.0509 | -3.7678
2.1182 | -34.051 | -3.7781
2.1243 | -34.0511 | -3.789
2.1304 | -34.0512 | -3.7992
2.1365 | -34.0515 | -3.8103
2.1426 | -34.0516 | -3.823
2.1487 | -34.0519 | -3.8346
2.1548 | -34.0522 | -3.847
2.1609 | -34.0526 | -3.8588
2.167 | -34.0529 | -3.8713
2.1731 | -34.0531 | -3.8825
2.1792 | -34.0538 | -3.8937
2.1853 | -34.0542 | -3.9059
2.1914 | -34.0546 | -3.9179
2.1975 | -34.055 | -3.9298
2.2036 | -34.0554 | -3.9423
2.2097 | -34.0557 | -3.954
2.2158 | -34.0566 | -3.9646
2.222 | -34.0571 | -3.9751
2.2281 | -34.0573 | -3.9857
2.2342 | -34.0577 | -3.9986
2.2403 | -34.0582 | -4.009
2.2464 | -34.0589 | -4.0204
2.2525 | -34.0593 | -4.0309
2.2586 | -34.0601 | -4.0411
2.2647 | -34.0608 | -4.0509
2.2708 | -34.0613 | -4.0625
2.2769 | -34.0621 | -4.0743
2.283 | -34.0625 | -4.0836
2.2891 | -34.0631 | -4.0941
2.2952 | -34.0637 | -4.1043
2.3013 | -34.0643 | -4.1146
2.3074 | -34.0649 | -4.1247
2.3135 | -34.0654 | -4.1358
2.3196 | -34.066 | -4.1451
2.3257 | -34.0663 | -4.156
2.3318 | -34.0667 | -4.165
2.3379 | -34.0672 | -4.1767
2.344 | -34.0675 | -4.1874
2.3501 | -34.068 | -4.1984
2.3562 | -34.0683 | -4.2073
2.3623 | -34.0686 | -4.2165
2.3685 | -34.0689 | -4.226
2.3746 | -34.0692 | -4.2364
2.3807 | -34.0697 | -4.2458
2.3868 | -34.0699 | -4.255
2.3929 | -34.0703 | -4.2647
2.399 | -34.0706 | -4.2731
2.4051 | -34.0711 | -4.283
2.4112 | -34.0714 | -4.2924
2.4173 | -34.0717 | -4.3016
2.4234 | -34.0718 | -4.3103
2.4295 | -34.0721 | -4.319
2.4356 | -34.0722 | -4.3286
2.4417 | -34.0723 | -4.3386
2.4478 | -34.0725 | -4.3484
2.4539 | -34.0728 | -4.3569
2.46 | -34.0726 | -4.3663
2.4661 | -34.0726 | -4.3755
2.4722 | -34.0729 | -4.3849
2.4783 | -34.0731 | -4.3937
2.4844 | -34.0732 | -4.4023
2.4905 | -34.0731 | -4.4116
2.4966 | -34.0729 | -4.4208
2.5027 | -34.0729 | -4.4292
2.5089 | -34.0729 | -4.4378
2.515 | -34.0728 | -4.4473
2.5211 | -34.0728 | -4.4558
2.5272 | -34.0725 | -4.4649
2.5333 | -34.0724 | -4.4745
2.5394 | -34.0724 | -4.4837
2.5455 | -34.0723 | -4.4928
2.5516 | -34.0722 | -4.5021
2.5577 | -34.0722 | -4.5117
2.5638 | -34.072 | -4.5214
2.5699 | -34.0719 | -4.5314
2.576 | -34.0718 | -4.5414
2.5821 | -34.0716 | -4.5511
2.5882 | -34.0715 | -4.5606
2.5943 | -34.0714 | -4.5705
2.6004 | -34.0712 | -4.5804
2.6065 | -34.0711 | -4.5893
2.6126 | -34.071 | -4.5987
2.6187 | -34.0709 | -4.6077
2.6248 | -34.0708 | -4.6169
2.6309 | -34.0707 | -4.6277
2.637 | -34.0707 | -4.6383
2.6431 | -34.0707 | -4.6481
2.6492 | -34.0706 | -4.6578
2.6554 | -34.0706 | -4.6681
2.6615 | -34.0705 | -4.6779
2.6676 | -34.0704 | -4.6878
2.6737 | -34.0705 | -4.6983
2.6798 | -34.0704 | -4.708
2.6859 | -34.0704 | -4.7181
2.692 | -34.0703 | -4.7279
2.6981 | -34.0703 | -4.7388
2.7042 | -34.0703 | -4.7505
2.7103 | -34.0703 | -4.7608
2.7164 | -34.0705 | -4.7714
2.7225 | -34.0705 | -4.7823
2.7286 | -34.0705 | -4.7933
2.7347 | -34.0705 | -4.8035
2.7408 | -34.0708 | -4.8132
2.7469 | -34.0707 | -4.8242
2.753 | -34.0709 | -4.835
2.7591 | -34.071 | -4.8465
2.7652 | -34.071 | -4.8576
2.7713 | -34.071 | -4.8675
2.7774 | -34.0714 | -4.8787
2.7835 | -34.0715 | -4.8892
2.7896 | -34.0716 | -4.8986
2.7958 | -34.0719 | -4.9089
2.8019 | -34.0724 | -4.9189
2.808 | -34.0726 | -4.9308
2.8141 | -34.0729 | -4.9434
2.8202 | -34.0732 | -4.9555
2.8263 | -34.0736 | -4.9663
2.8324 | -34.0737 | -4.9782
2.8385 | -34.074 | -4.9901
2.8446 | -34.0744 | -5.0017
2.8507 | -34.0746 | -5.0138
2.8568 | -34.075 | -5.025
2.8629 | -34.0753 | -5.0365
2.869 | -34.0757 | -5.0475
2.8751 | -34.076 | -5.0594
2.8812 | -34.0765 | -5.0722
2.8873 | -34.0767 | -5.0834
2.8934 | -34.077 | -5.095
2.8995 | -34.0772 | -5.1067
2.9056 | -34.0775 | -5.1171
2.9117 | -34.0779 | -5.1284
2.9178 | -34.0784 | -5.141
2.9239 | -34.0787 | -5.1537
2.93 | -34.079 | -5.1665
2.9361 | -34.0795 | -5.1779
2.9423 | -34.0799 | -5.19
2.9484 | -34.0803 | -5.2032
2.9545 | -34.0808 | -5.2162
2.9606 | -34.0812 | -5.229
2.9667 | -34.0816 | -5.2406
2.9728 | -34.0819 | -5.2523
2.9789 | -34.0821 | -5.2647
2.985 | -34.0824 | -5.278
2.9911 | -34.0828 | -5.2902
2.9972 | -34.0831 | -5.3025
3.0033 | -34.0835 | -5.315
3.0094 | -34.0839 | -5.3277
3.0155 | -34.0844 | -5.3402
3.0216 | -34.0847 | -5.3532
3.0277 | -34.0851 | -5.3662
3.0338 | -34.0855 | -5.3792
3.0399 | -34.0859 | -5.3925
3.046 | -34.0862 | -5.4049
3.0521 | -34.0866 | -5.417
3.0582 | -34.0869 | -5.43
3.0643 | -34.0873 | -5.4425
3.0704 | -34.0875 | -5.455
3.0765 | -34.0878 | -5.467
3.0827 | -34.088 | -5.4798
3.0888 | -34.0882 | -5.4923
3.0949 | -34.0884 | -5.5048
3.101 | -34.0888 | -5.5184
3.1071 | -34.089 | -5.5318
3.1132 | -34.0893 | -5.5451
3.1193 | -34.0896 | -5.5576
3.1254 | -34.0899 | -5.5693
3.1315 | -34.0902 | -5.5827
3.1376 | -34.0903 | -5.595
3.1437 | -34.0906 | -5.608
3.1498 | -34.0908 | -5.6213
3.1559 | -34.0911 | -5.634
3.162 | -34.0911 | -5.6472
3.1681 | -34.0911 | -5.6605
3.1742 | -34.0914 | -5.6736
3.1803 | -34.0917 | -5.6862
3.1864 | -34.0919 | -5.698
3.1925 | -34.0922 | -5.7111
3.1986 | -34.0922 | -5.723
3.2047 | -34.0923 | -5.7361
3.2108 | -34.0925 | -5.7476
3.2169 | -34.0928 | -5.7606
3.223 | -34.0928 | -5.7734
3.2292 | -34.0927 | -5.7865
3.2353 | -34.0926 | -5.796
3.2414 | -34.0927 | -5.8085
3.2475 | -34.0927 | -5.8191
3.2536 | -34.0929 | -5.8323
3.2597 | -34.0931 | -5.8434
3.2658 | -34.0932 | -5.854
3.2719 | -34.0934 | -5.8649
3.278 | -34.0934 | -5.8765
3.2841 | -34.0935 | -5.8882
3.2902 | -34.0935 | -5.9002
3.2963 | -34.0936 | -5.9114
3.3024 | -34.0936 | -5.9214
3.3085 | -34.0938 | -5.9333
3.3146 | -34.0937 | -5.944
3.3207 | -34.0939 | -5.9556
3.3268 | -34.094 | -5.9673
3.3329 | -34.0941 | -5.9771
3.339 | -34.0942 | -5.9867
3.3451 | -34.0942 | -5.9978
3.3512 | -34.0944 | -6.0073
3.3573 | -34.0945 | -6.0183
3.3634 | -34.0946 | -6.0296
3.3696 | -34.0946 | -6.041
3.3757 | -34.0946 | -6.0508
3.3818 | -34.0949 | -6.0624
3.3879 | -34.0948 | -6.0726
3.394 | -34.0948 | -6.0826
3.4001 | -34.0949 | -6.0934
3.4062 | -34.0949 | -6.1035
3.4123 | -34.095 | -6.1134
3.4184 | -34.0951 | -6.124
3.4245 | -34.0951 | -6.1345
3.4306 | -34.0951 | -6.1443
3.4367 | -34.0953 | -6.1541
3.4428 | -34.0956 | -6.1635
3.4489 | -34.0956 | -6.1745
3.455 | -34.0955 | -6.1848
3.4611 | -34.0956 | -6.1956
3.4672 | -34.0956 | -6.2066
3.4733 | -34.0957 | -6.2165
3.4794 | -34.0957 | -6.2268
3.4855 | -34.096 | -6.2366
3.4916 | -34.0961 | -6.2466
3.4977 | -34.0961 | -6.2571
3.5038 | -34.0963 | -6.267
3.5099 | -34.0964 | -6.2769
3.5161 | -34.0965 | -6.2866
3.5222 | -34.0967 | -6.2968
3.5283 | -34.0967 | -6.3068
3.5344 | -34.0968 | -6.3175
3.5405 | -34.0969 | -6.3275
3.5466 | -34.0971 | -6.3371
3.5527 | -34.0972 | -6.3476
3.5588 | -34.0974 | -6.3574
3.5649 | -34.0975 | -6.3676
3.571 | -34.0977 | -6.3779
3.5771 | -34.0978 | -6.3883
3.5832 | -34.098 | -6.3985
3.5893 | -34.0982 | -6.4083
3.5954 | -34.0983 | -6.4187
3.6015 | -34.0985 | -6.4289
3.6076 | -34.0987 | -6.4385
3.6137 | -34.0988 | -6.4488
3.6198 | -34.0991 | -6.458
3.6259 | -34.0993 | -6.4685
3.632 | -34.0995 | -6.4791
3.6381 | -34.0997 | -6.4888
3.6442 | -34.1 | -6.4993
3.6503 | -34.1002 | -6.5087
3.6565 | -34.1005 | -6.5202
3.6626 | -34.1006 | -6.531
3.6687 | -34.1009 | -6.5415
3.6748 | -34.1012 | -6.5515
3.6809 | -34.1015 | -6.5619
3.687 | -34.1018 | -6.5728
3.6931 | -34.1021 | -6.583
3.6992 | -34.1023 | -6.5935
3.7053 | -34.1026 | -6.6037
3.7114 | -34.1028 | -6.6149
3.7175 | -34.1031 | -6.6246
3.7236 | -34.1033 | -6.6337
3.7297 | -34.1034 | -6.6443
3.7358 | -34.1037 | -6.6557
3.7419 | -34.1039 | -6.6671
3.748 | -34.104 | -6.6778
3.7541 | -34.1042 | -6.6888
3.7602 | -34.1045 | -6.6995
3.7663 | -34.1047 | -6.7097
3.7724 | -34.105 | -6.7206
3.7785 | -34.1053 | -6.7311
3.7846 | -34.1055 | -6.7417
3.7907 | -34.1059 | -6.7524
3.7969 | -34.1061 | -6.7637
3.803 | -34.1064 | -6.7739
3.8091 | -34.1067 | -6.7842
3.8152 | -34.1069 | -6.7953
3.8213 | -34.1071 | -6.8067
3.8274 | -34.1075 | -6.8172
3.8335 | -34.1078 | -6.8285
3.8396 | -34.1082 | -6.8399
3.8457 | -34.1085 | -6.8499
3.8518 | -34.1089 | -6.8607
3.8579 | -34.1092 | -6.8711
3.864 | -34.1095 | -6.8816
3.8701 | -34.1098 | -6.8921
3.8762 | -34.1101 | -6.9033
3.8823 | -34.1104 | -6.9137
3.8884 | -34.1108 | -6.9238
3.8945 | -34.1111 | -6.9348
3.9006 | -34.1115 | -6.9459
3.9067 | -34.1118 | -6.9572
3.9128 | -34.1122 | -6.9683
3.9189 | -34.1125 | -6.9789
3.925 | -34.1128 | -6.9895
3.9311 | -34.1131 | -6.9996
3.9372 | -34.1135 | -7.0108
3.9434 | -34.1138 | -7.0211
3.9495 | -34.1141 | -7.0324
3.9556 | -34.1144 | -7.0435
3.9617 | -34.1147 | -7.0545
3.9678 | -34.115 | -7.0651
3.9739 | -34.1154 | -7.0761
3.98 | -34.1158 | -7.0877
3.9861 | -34.116 | -7.098
3.9922 | -34.1163 | -7.1095
3.9983 | -34.1166 | -7.1196
4.0044 | -34.1169 | -7.1294
4.0105 | -34.1172 | -7.1401
4.0166 | -34.1176 | -7.1509
4.0227 | -34.118 | -7.161
4.0288 | -34.1182 | -7.1719
4.0349 | -34.1185 | -7.1818
4.041 | -34.1188 | -7.1926
4.0471 | -34.1191 | -7.2026
4.0532 | -34.1195 | -7.2138
4.0593 | -34.1197 | -7.2251
4.0654 | -34.12 | -7.2358
4.0715 | -34.1203 | -7.2467
4.0776 | -34.1206 | -7.2568
4.0838 | -34.1209 | -7.2677
4.0899 | -34.1212 | -7.279
4.096 | -34.1215 | -7.2892
4.1021 | -34.1218 | -7.3
4.1082 | -34.1222 | -7.3103
4.1143 | -34.1225 | -7.3203
4.1204 | -34.1229 | -7.3311
4.1265 | -34.1232 | -7.3415
4.1326 | -34.1236 | -7.3519
4.1387 | -34.1239 | -7.363
4.1448 | -34.1243 | -7.3737
4.1509 | -34.1245 | -7.3843
4.157 | -34.1248 | -7.3951
4.1631 | -34.1251 | -7.406
4.1692 | -34.1254 | -7.4167
4.1753 | -34.1257 | -7.4274
4.1814 | -34.126 | -7.4374
4.1875 | -34.1263 | -7.4485
4.1936 | -34.1266 | -7.4596
4.1997 | -34.1269 | -7.4699
4.2058 | -34.1271 | -7.4811
4.2119 | -34.1274 | -7.4908
4.218 | -34.1277 | -7.5011
4.2241 | -34.128 | -7.5107
4.2303 | -34.1284 | -7.522
4.2364 | -34.1288 | -7.5333
4.2425 | -34.1291 | -7.5438
4.2486 | -34.1293 | -7.5547
4.2547 | -34.1295 | -7.5641
4.2608 | -34.1298 | -7.5739
4.2669 | -34.1301 | -7.5845
4.273 | -34.1304 | -7.5951
4.2791 | -34.1306 | -7.6057
4.2852 | -34.1308 | -7.6158
4.2913 | -34.131 | -7.6253
4.2974 | -34.1313 | -7.6359
4.3035 | -34.1315 | -7.6465
4.3096 | -34.1317 | -7.6568
4.3157 | -34.1319 | -7.6668
4.3218 | -34.1322 | -7.6768
4.3279 | -34.1324 | -7.6872
4.334 | -34.1327 | -7.6971
4.3401 | -34.1329 | -7.7073
4.3462 | -34.1332 | -7.7179
4.3523 | -34.1335 | -7.7279
4.3584 | -34.1337 | -7.7367
4.3645 | -34.1339 | -7.7463
4.3707 | -34.1341 | -7.7559
4.3768 | -34.1343 | -7.7663
4.3829 | -34.1345 | -7.7764
4.389 | -34.1348 | -7.7861
4.3951 | -34.1351 | -7.7958
4.4012 | -34.1353 | -7.8054
4.4073 | -34.1355 | -7.8157
4.4134 | -34.1358 | -7.826
4.4195 | -34.136 | -7.8356
4.4256 | -34.1363 | -7.8454
4.4317 | -34.1365 | -7.855
4.4378 | -34.1367 | -7.8649
4.4439 | -34.1369 | -7.8746
4.45 | -34.1371 | -7.8847
4.4561 | -34.1372 | -7.8947
4.4622 | -34.1374 | -7.9043
4.4683 | -34.1376 | -7.9136
4.4744 | -34.1378 | -7.9237
4.4805 | -34.1381 | -7.9331
4.4866 | -34.1383 | -7.9424
4.4927 | -34.1385 | -7.9525
4.4988 | -34.1387 | -7.9621
4.5049 | -34.1389 | -7.9716
4.511 | -34.1392 | -7.9809
4.5172 | -34.1395 | -7.9908
4.5233 | -34.1397 | -8.0012
4.5294 | -34.1399 | -8.0107
4.5355 | -34.1402 | -8.0202
4.5416 | -34.1405 | -8.0295
4.5477 | -34.1406 | -8.0388
4.5538 | -34.1408 | -8.048
4.5599 | -34.1412 | -8.0576
4.566 | -34.1414 | -8.0674
4.5721 | -34.1416 | -8.0767
4.5782 | -34.1419 | -8.0864
4.5843 | -34.1421 | -8.0961
4.5904 | -34.1424 | -8.1054
4.5965 | -34.1427 | -8.1143
4.6026 | -34.143 | -8.1248
4.6087 | -34.1432 | -8.134
4.6148 | -34.1434 | -8.1445
4.6209 | -34.1438 | -8.1531
4.627 | -34.1441 | -8.163
4.6331 | -34.1442 | -8.1721
4.6392 | -34.1445 | -8.1814
4.6453 | -34.1447 | -8.1909
4.6514 | -34.145 | -8.2004
4.6576 | -34.1451 | -8.2107
4.6637 | -34.1453 | -8.2187
4.6698 | -34.1456 | -8.2275
4.6759 | -34.1457 | -8.2371
4.682 | -34.1459 | -8.2464
4.6881 | -34.1462 | -8.2553
4.6942 | -34.1464 | -8.2647
4.7003 | -34.1466 | -8.2742
4.7064 | -34.1469 | -8.2836
4.7125 | -34.1472 | -8.2929
4.7186 | -34.1474 | -8.3015
4.7247 | -34.1477 | -8.3111
4.7308 | -34.148 | -8.3205
4.7369 | -34.1482 | -8.3301
4.743 | -34.1484 | -8.3402
4.7491 | -34.1486 | -8.3496
4.7552 | -34.1487 | -8.3588
4.7613 | -34.149 | -8.3684
4.7674 | -34.1493 | -8.3771
4.7735 | -34.1494 | -8.387
4.7796 | -34.1496 | -8.3966
4.7857 | -34.1497 | -8.4062
4.7918 | -34.15 | -8.4149
4.7979 | -34.1503 | -8.424
4.8041 | -34.1504 | -8.4333
4.8102 | -34.1507 | -8.4426
4.8163 | -34.1509 | -8.4521
4.8224 | -34.151 | -8.4608
4.8285 | -34.1514 | -8.47
4.8346 | -34.1516 | -8.4796
4.8407 | -34.1518 | -8.4889
4.8468 | -34.152 | -8.4984
4.8529 | -34.1523 | -8.5079
4.859 | -34.1525 | -8.5172
4.8651 | -34.1527 | -8.5267
4.8712 | -34.1529 | -8.5362
4.8773 | -34.1532 | -8.5457
4.8834 | -34.1534 | -8.5554
4.8895 | -34.1537 | -8.5655
4.8956 | -34.1539 | -8.5755
4.9017 | -34.1541 | -8.5851
4.9078 | -34.1543 | -8.5947
4.9139 | -34.1545 | -8.6035
4.92 | -34.1548 | -8.6126
4.9261 | -34.155 | -8.6217
4.9322 | -34.1554 | -8.6308
4.9383 | -34.1557 | -8.6398
4.9445 | -34.156 | -8.6499
4.9506 | -34.1562 | -8.6589
4.9567 | -34.1565 | -8.6683
4.9628 | -34.1567 | -8.6776
4.9689 | -34.1569 | -8.6871
4.975 | -34.1571 | -8.6963
4.9811 | -34.1574 | -8.7062
4.9872 | -34.1576 | -8.7161
4.9933 | -34.1579 | -8.7254
4.9994 | -34.1582 | -8.7349
5.0055 | -34.1586 | -8.7445
5.0116 | -34.1589 | -8.7544
5.0177 | -34.1592 | -8.7642
5.0238 | -34.1595 | -8.7744
5.0299 | -34.1599 | -8.785
5.036 | -34.1601 | -8.7945
5.0421 | -34.1603 | -8.8042
5.0482 | -34.1606 | -8.8144
5.0543 | -34.1609 | -8.8237
5.0604 | -34.1612 | -8.8339
5.0665 | -34.1615 | -8.8436
5.0726 | -34.1618 | -8.8537
5.0787 | -34.1621 | -8.8637
5.0848 | -34.1624 | -8.8736
5.091 | -34.1627 | -8.8835
5.0971 | -34.163 | -8.8937
5.1032 | -34.1633 | -8.9037
5.1093 | -34.1636 | -8.9135
5.1154 | -34.1638 | -8.9236
5.1215 | -34.1641 | -8.9337
5.1276 | -34.1644 | -8.9435
5.1337 | -34.1648 | -8.9532
5.1398 | -34.165 | -8.9631
5.1459 | -34.1653 | -8.9733
5.152 | -34.1656 | -8.9841
5.1581 | -34.1658 | -8.9946
5.1642 | -34.1661 | -9.0044
5.1703 | -34.1664 | -9.0143
5.1764 | -34.1667 | -9.0242
5.1825 | -34.167 | -9.0347
5.1886 | -34.1672 | -9.0446
5.1947 | -34.1675 | -9.0553
5.2008 | -34.1679 | -9.065
5.2069 | -34.1682 | -9.0759
5.213 | -34.1685 | -9.0855
5.2191 | -34.1688 | -9.0957
5.2252 | -34.1691 | -9.1058
5.2314 | -34.1693 | -9.1163
5.2375 | -34.1697 | -9.1263
5.2436 | -34.17 | -9.1364
5.2497 | -34.1703 | -9.1464
5.2558 | -34.1706 | -9.157
5.2619 | -34.1709 | -9.1676
5.268 | -34.1712 | -9.1777
5.2741 | -34.1715 | -9.1887
5.2802 | -34.1718 | -9.1992
5.2863 | -34.1722 | -9.2093
5.2924 | -34.1724 | -9.2202
5.2985 | -34.1728 | -9.2314
5.3046 | -34.1732 | -9.2425
5.3107 | -34.1736 | -9.253
5.3168 | -34.174 | -9.2633
5.3229 | -34.1743 | -9.273
5.329 | -34.1746 | -9.2828
5.3351 | -34.175 | -9.2931
5.3412 | -34.1754 | -9.3031
5.3473 | -34.1758 | -9.3139
5.3534 | -34.1761 | -9.3245
5.3595 | -34.1764 | -9.3351
5.3656 | -34.1767 | -9.3459
5.3717 | -34.177 | -9.3561
5.3779 | -34.1774 | -9.3666
5.384 | -34.1776 | -9.3772
5.3901 | -34.178 | -9.3876
5.3962 | -34.1782 | -9.3984
5.4023 | -34.1786 | -9.4088
5.4084 | -34.1789 | -9.4191
5.4145 | -34.1792 | -9.4301
5.4206 | -34.1795 | -9.4412
5.4267 | -34.1798 | -9.4522
5.4328 | -34.1802 | -9.4621
5.4389 | -34.1805 | -9.4724
5.445 | -34.1808 | -9.4834
5.4511 | -34.1811 | -9.4937
5.4572 | -34.1815 | -9.5046
5.4633 | -34.1818 | -9.5148
5.4694 | -34.1821 | -9.5252
5.4755 | -34.1824 | -9.5361
5.4816 | -34.1827 | -9.5465
5.4877 | -34.183 | -9.5571
5.4938 | -34.1834 | -9.5676
5.4999 | -34.1837 | -9.5782
5.506 | -34.184 | -9.5887
5.5121 | -34.1843 | -9.5998
5.5183 | -34.1846 | -9.6099
5.5244 | -34.1848 | -9.6204
5.5305 | -34.1852 | -9.6311
5.5366 | -34.1855 | -9.6416
5.5427 | -34.1858 | -9.6527
5.5488 | -34.1861 | -9.6635
5.5549 | -34.1864 | -9.6741
5.561 | -34.1867 | -9.6851
5.5671 | -34.1871 | -9.6954
5.5732 | -34.1873 | -9.7061
5.5793 | -34.1877 | -9.7172
5.5854 | -34.1881 | -9.728
5.5915 | -34.1883 | -9.7386
5.5976 | -34.1886 | -9.7494
5.6037 | -34.1889 | -9.7599
5.6098 | -34.1892 | -9.7703
5.6159 | -34.1895 | -9.7812
5.622 | -34.1899 | -9.7914
5.6281 | -34.1902 | -9.8021
5.6342 | -34.1905 | -9.8126
5.6403 | -34.1908 | -9.8226
5.6464 | -34.1912 | -9.8332
5.6525 | -34.1914 | -9.8434
5.6586 | -34.1918 | -9.8534
5.6648 | -34.192 | -9.8645
5.6709 | -34.1923 | -9.8753
5.677 | -34.1927 | -9.8866
5.6831 | -34.193 | -9.8966
5.6892 | -34.1933 | -9.9071
5.6953 | -34.1936 | -9.9176
5.7014 | -34.1939 | -9.9287
5.7075 | -34.1942 | -9.9395
5.7136 | -34.1945 | -9.9499
5.7197 | -34.1948 | -9.9604
5.7258 | -34.1951 | -9.9709
5.7319 | -34.1954 | -9.9814
5.738 | -34.1958 | -9.992
5.7441 | -34.1961 | -10.0025
5.7502 | -34.1964 | -10.0129
5.7563 | -34.1967 | -10.0231
5.7624 | -34.197 | -10.0342
5.7685 | -34.1972 | -10.0445
5.7746 | -34.1975 | -10.0548
5.7807 | -34.1978 | -10.0649
5.7868 | -34.1981 | -10.0753
5.7929 | -34.1984 | -10.0857
5.799 | -34.1987 | -10.0962
5.8052 | -34.199 | -10.1069
5.8113 | -34.1992 | -10.1176
5.8174 | -34.1995 | -10.1283
5.8235 | -34.1998 | -10.1387
5.8296 | -34.2001 | -10.1496
5.8357 | -34.2003 | -10.1602
5.8418 | -34.2007 | -10.1708
5.8479 | -34.2009 | -10.1816
5.854 | -34.2012 | -10.1922
5.8601 | -34.2015 | -10.2026
5.8662 | -34.2018 | -10.2133
5.8723 | -34.202 | -10.2233
5.8784 | -34.2023 | -10.2338
5.8845 | -34.2027 | -10.2445
5.8906 | -34.2029 | -10.2549
5.8967 | -34.2032 | -10.2651
5.9028 | -34.2035 | -10.2757
5.9089 | -34.2037 | -10.2856
5.915 | -34.2041 | -10.2962
5.9211 | -34.2044 | -10.3067
5.9272 | -34.2046 | -10.3173
5.9333 | -34.2049 | -10.3271
5.9394 | -34.2053 | -10.3374
5.9455 | -34.2055 | -10.3475
5.9517 | -34.2059 | -10.358
5.9578 | -34.2062 | -10.3686
5.9639 | -34.2065 | -10.3789
5.97 | -34.2068 | -10.3894
5.9761 | -34.2071 | -10.3995
5.9822 | -34.2074 | -10.41
5.9883 | -34.2077 | -10.4201
5.9944 | -34.208 | -10.4305
6.0005 | -34.2084 | -10.441
6.0066 | -34.2086 | -10.4511
6.0127 | -34.2089 | -10.461
6.0188 | -34.2092 | -10.4715
6.0249 | -34.2095 | -10.4818
6.031 | -34.2098 | -10.4921
6.0371 | -34.2102 | -10.5027
6.0432 | -34.2105 | -10.5131
6.0493 | -34.2109 | -10.5233
6.0554 | -34.2112 | -10.5337
6.0615 | -34.2115 | -10.5445
6.0676 | -34.2118 | -10.5553
6.0737 | -34.2121 | -10.5658
6.0798 | -34.2124 | -10.576
6.0859 | -34.2127 | -10.5862
6.0921 | -34.213 | -10.5965
6.0982 | -34.2134 | -10.6071
6.1043 | -34.2137 | -10.6173
6.1104 | -34.2141 | -10.628
6.1165 | -34.2144 | -10.6378
6.1226 | -34.2147 | -10.6476
6.1287 | -34.2151 | -10.658
6.1348 | -34.2154 | -10.6684
6.1409 | -34.2157 | -10.6786
6.147 | -34.2161 | -10.6888
6.1531 | -34.2164 | -10.6991
6.1592 | -34.2167 | -10.7095
6.1653 | -34.217 | -10.7196
6.1714 | -34.2174 | -10.7299
6.1775 | -34.2177 | -10.7404
6.1836 | -34.218 | -10.7505
6.1897 | -34.2183 | -10.7607
6.1958 | -34.2186 | -10.7711
6.2019 | -34.2189 | -10.7815
6.208 | -34.2192 | -10.7922
6.2141 | -34.2196 | -10.8027
6.2202 | -34.2199 | -10.8134
6.2263 | -34.2203 | -10.8235
6.2325 | -34.2206 | -10.8335
6.2386 | -34.221 | -10.8437
6.2447 | -34.2213 | -10.854
6.2508 | -34.2216 | -10.8642
6.2569 | -34.2219 | -10.8748
6.263 | -34.2223 | -10.8847
6.2691 | -34.2227 | -10.895
6.2752 | -34.223 | -10.9052
6.2813 | -34.2234 | -10.9155
6.2874 | -34.2237 | -10.9258
6.2935 | -34.2241 | -10.9359
6.2996 | -34.2244 | -10.9463
6.3057 | -34.2247 | -10.9564
6.3118 | -34.2251 | -10.9665
6.3179 | -34.2255 | -10.9768
6.324 | -34.2258 | -10.9872
6.3301 | -34.2262 | -10.9972
6.3362 | -34.2265 | -11.0076
6.3423 | -34.2269 | -11.0177
6.3484 | -34.2273 | -11.0281
6.3545 | -34.2277 | -11.0379
6.3606 | -34.228 | -11.0481
6.3667 | -34.2285 | -11.0582
6.3728 | -34.2288 | -11.0684
6.379 | -34.2291 | -11.0786
6.3851 | -34.2295 | -11.0891
6.3912 | -34.2299 | -11.0991
6.3973 | -34.2303 | -11.109
6.4034 | -34.2306 | -11.1192
6.4095 | -34.2309 | -11.1297
6.4156 | -34.2313 | -11.1396
6.4217 | -34.2316 | -11.1498
6.4278 | -34.232 | -11.1599
6.4339 | -34.2324 | -11.17
6.44 | -34.2327 | -11.1799
6.4461 | -34.2331 | -11.1905
6.4522 | -34.2335 | -11.2002
6.4583 | -34.2338 | -11.2105
6.4644 | -34.2341 | -11.221
6.4705 | -34.2345 | -11.2315
6.4766 | -34.2349 | -11.2421
6.4827 | -34.2352 | -11.2527
6.4888 | -34.2355 | -11.2629
6.4949 | -34.2359 | -11.2729
6.501 | -34.2362 | -11.2833
6.5071 | -34.2366 | -11.2937
6.5132 | -34.237 | -11.3037
6.5194 | -34.2374 | -11.3138
6.5255 | -34.2378 | -11.3237
6.5316 | -34.2381 | -11.334
6.5377 | -34.2385 | -11.3445
6.5438 | -34.2388 | -11.3547
6.5499 | -34.2392 | -11.3648
6.556 | -34.2396 | -11.3749
6.5621 | -34.2399 | -11.3851
6.5682 | -34.2402 | -11.395
6.5743 | -34.2406 | -11.406
6.5804 | -34.2409 | -11.4161
6.5865 | -34.2412 | -11.4265
6.5926 | -34.2416 | -11.4367
6.5987 | -34.242 | -11.4467
6.6048 | -34.2423 | -11.4575
6.6109 | -34.2427 | -11.4676
6.617 | -34.243 | -11.4786
6.6231 | -34.2433 | -11.489
6.6292 | -34.2438 | -11.4991
6.6353 | -34.2441 | -11.5101
6.6414 | -34.2445 | -11.5206
6.6475 | -34.2449 | -11.5309
6.6536 | -34.2453 | -11.5408
6.6597 | -34.2457 | -11.551
6.6659 | -34.246 | -11.561
6.672 | -34.2464 | -11.5714
6.6781 | -34.2468 | -11.5814
6.6842 | -34.2471 | -11.5918
6.6903 | -34.2475 | -11.6019
6.6964 | -34.2479 | -11.6122
6.7025 | -34.2482 | -11.6226
6.7086 | -34.2485 | -11.6326
6.7147 | -34.2489 | -11.6429
6.7208 | -34.2493 | -11.6536
6.7269 | -34.2496 | -11.6634
6.733 | -34.2499 | -11.6731
6.7391 | -34.2503 | -11.6834
6.7452 | -34.2508 | -11.6932
6.7513 | -34.2511 | -11.7032
6.7574 | -34.2515 | -11.7136
6.7635 | -34.2519 | -11.7234
6.7696 | -34.2523 | -11.7338
6.7757 | -34.2527 | -11.7442
6.7818 | -34.2531 | -11.7545
6.7879 | -34.2534 | -11.7648
6.794 | -34.2538 | -11.775
6.8001 | -34.2542 | -11.785
6.8063 | -34.2545 | -11.7954
6.8124 | -34.2549 | -11.8055
6.8185 | -34.2552 | -11.8156
6.8246 | -34.2556 | -11.8259
6.8307 | -34.256 | -11.836
6.8368 | -34.2564 | -11.8463
6.8429 | -34.2568 | -11.8563
6.849 | -34.2571 | -11.8665
6.8551 | -34.2575 | -11.8767
6.8612 | -34.2579 | -11.8868
6.8673 | -34.2582 | -11.8968
6.8734 | -34.2586 | -11.9076
6.8795 | -34.259 | -11.918
6.8856 | -34.2594 | -11.9277
6.8917 | -34.2598 | -11.9373
6.8978 | -34.2601 | -11.9476
6.9039 | -34.2604 | -11.9578
6.91 | -34.2608 | -11.9679
6.9161 | -34.2612 | -11.978
6.9222 | -34.2615 | -11.9883
6.9283 | -34.2618 | -11.9982
6.9344 | -34.2622 | -12.0081
6.9405 | -34.2625 | -12.0181
6.9466 | -34.2629 | -12.0283
6.9528 | -34.2632 | -12.0383
6.9589 | -34.2635 | -12.0482
6.965 | -34.2639 | -12.0582
6.9711 | -34.2643 | -12.0682
6.9772 | -34.2647 | -12.0781
6.9833 | -34.265 | -12.0885
6.9894 | -34.2653 | -12.0988
6.9955 | -34.2657 | -12.109
7.0016 | -34.266 | -12.1195
7.0077 | -34.2664 | -12.1299
7.0138 | -34.2668 | -12.1404
7.0199 | -34.2671 | -12.1504
7.026 | -34.2674 | -12.1604
7.0321 | -34.2677 | -12.1707
7.0382 | -34.268 | -12.1807
7.0443 | -34.2684 | -12.191
7.0504 | -34.2687 | -12.2014
7.0565 | -34.2691 | -12.2117
7.0626 | -34.2694 | -12.2219
7.0687 | -34.2697 | -12.2318
7.0748 | -34.2701 | -12.2419
7.0809 | -34.2704 | -12.2521
7.087 | -34.2708 | -12.2622
7.0932 | -34.2711 | -12.2723
7.0993 | -34.2714 | -12.2825
7.1054 | -34.2718 | -12.2924
7.1115 | -34.272 | -12.3025
7.1176 | -34.2723 | -12.3125
7.1237 | -34.2726 | -12.3226
7.1298 | -34.2729 | -12.3323
7.1359 | -34.2732 | -12.3421
7.142 | -34.2735 | -12.3524
7.1481 | -34.2739 | -12.3623
7.1542 | -34.2742 | -12.3722
7.1603 | -34.2745 | -12.3821
7.1664 | -34.2749 | -12.3922
7.1725 | -34.2752 | -12.402
7.1786 | -34.2755 | -12.4121
7.1847 | -34.2759 | -12.422
7.1908 | -34.2762 | -12.432
7.1969 | -34.2765 | -12.4419
7.203 | -34.2768 | -12.4521
7.2091 | -34.2772 | -12.4619
7.2152 | -34.2775 | -12.4719
7.2213 | -34.2778 | -12.4819
7.2274 | -34.2781 | -12.4919
7.2335 | -34.2784 | -12.5019
7.2397 | -34.2787 | -12.5119
7.2458 | -34.279 | -12.5216
7.2519 | -34.2793 | -12.5315
7.258 | -34.2797 | -12.5416
7.2641 | -34.28 | -12.5515
7.2702 | -34.2804 | -12.5617
7.2763 | -34.2807 | -12.5714
7.2824 | -34.2811 | -12.5811
7.2885 | -34.2814 | -12.591
7.2946 | -34.2817 | -12.6012
7.3007 | -34.282 | -12.6111
7.3068 | -34.2823 | -12.6212
7.3129 | -34.2827 | -12.6308
7.319 | -34.283 | -12.6409
7.3251 | -34.2833 | -12.6506
7.3312 | -34.2837 | -12.6602
7.3373 | -34.284 | -12.67
7.3434 | -34.2844 | -12.6797
7.3495 | -34.2847 | -12.6892
7.3556 | -34.285 | -12.6994
7.3617 | -34.2853 | -12.7096
7.3678 | -34.2857 | -12.7197
7.3739 | -34.286 | -12.7291
7.3801 | -34.2863 | -12.7387
7.3862 | -34.2866 | -12.7485
7.3923 | -34.2869 | -12.7582
7.3984 | -34.2873 | -12.7682
7.4045 | -34.2876 | -12.7779
7.4106 | -34.2879 | -12.7876
7.4167 | -34.2882 | -12.7977
7.4228 | -34.2885 | -12.8073
7.4289 | -34.2888 | -12.8169
7.435 | -34.2892 | -12.8266
7.4411 | -34.2895 | -12.8362
7.4472 | -34.2899 | -12.8461
7.4533 | -34.2902 | -12.8561
7.4594 | -34.2905 | -12.8661
7.4655 | -34.2909 | -12.8759
7.4716 | -34.2912 | -12.8858
7.4777 | -34.2915 | -12.8952
7.4838 | -34.2919 | -12.9053
7.4899 | -34.2923 | -12.9152
7.496 | -34.2926 | -12.9253
7.5021 | -34.2929 | -12.9349
7.5082 | -34.2932 | -12.9448
7.5143 | -34.2935 | -12.9545
7.5204 | -34.2939 | -12.9639
7.5266 | -34.2941 | -12.9739
7.5327 | -34.2945 | -12.9836
7.5388 | -34.2948 | -12.9931
7.5449 | -34.2951 | -13.003
7.551 | -34.2954 | -13.0131
7.5571 | -34.2957 | -13.0227
7.5632 | -34.296 | -13.0323
7.5693 | -34.2963 | -13.0423
7.5754 | -34.2967 | -13.0522
7.5815 | -34.2971 | -13.0616
7.5876 | -34.2973 | -13.0711
7.5937 | -34.2976 | -13.0806
7.5998 | -34.2979 | -13.0904
7.6059 | -34.2982 | -13.1
7.612 | -34.2985 | -13.1097
7.6181 | -34.2988 | -13.1194
7.6242 | -34.2992 | -13.1291
7.6303 | -34.2996 | -13.1384
7.6364 | -34.2999 | -13.1483
7.6425 | -34.3002 | -13.1583
7.6486 | -34.3005 | -13.1681
7.6547 | -34.3009 | -13.178
7.6608 | -34.3012 | -13.1878
7.667 | -34.3015 | -13.1969
7.6731 | -34.3018 | -13.2067
7.6792 | -34.3021 | -13.2164
7.6853 | -34.3024 | -13.2259
7.6914 | -34.3026 | -13.2357
7.6975 | -34.3029 | -13.2456
7.7036 | -34.3032 | -13.2554
7.7097 | -34.3036 | -13.2649
7.7158 | -34.3039 | -13.2744
7.7219 | -34.3042 | -13.2842
7.728 | -34.3044 | -13.294
7.7341 | -34.3047 | -13.3037
7.7402 | -34.3051 | -13.3132
7.7463 | -34.3054 | -13.3229
7.7524 | -34.3057 | -13.3327
7.7585 | -34.306 | -13.3426
7.7646 | -34.3063 | -13.352
7.7707 | -34.3066 | -13.3614
7.7768 | -34.3068 | -13.3709
7.7829 | -34.3071 | -13.3804
7.789 | -34.3074 | -13.3897
7.7951 | -34.3077 | -13.3997
7.8012 | -34.308 | -13.4098
7.8073 | -34.3083 | -13.4191
7.8135 | -34.3086 | -13.4289
7.8196 | -34.3089 | -13.4383
7.8257 | -34.3092 | -13.4483
7.8318 | -34.3095 | -13.4583
7.8379 | -34.3098 | -13.4676
7.844 | -34.3101 | -13.4768
7.8501 | -34.3103 | -13.4866
7.8562 | -34.3106 | -13.4968
7.8623 | -34.3109 | -13.5065
7.8684 | -34.3112 | -13.5158
7.8745 | -34.3115 | -13.5253
7.8806 | -34.3118 | -13.5344
7.8867 | -34.3121 | -13.544
7.8928 | -34.3124 | -13.5536
7.8989 | -34.3128 | -13.5627
7.905 | -34.3131 | -13.5727
7.9111 | -34.3134 | -13.5823
7.9172 | -34.3137 | -13.5917
7.9233 | -34.3139 | -13.6013
7.9294 | -34.3143 | -13.611
7.9355 | -34.3146 | -13.6203
7.9416 | -34.3148 | -13.6299
7.9477 | -34.3152 | -13.6391
7.9539 | -34.3155 | -13.6485
7.96 | -34.3158 | -13.6583
7.9661 | -34.3161 | -13.668
7.9722 | -34.3164 | -13.6778
7.9783 | -34.3167 | -13.6872
7.9844 | -34.317 | -13.6972
7.9905 | -34.3173 | -13.7065
7.9966 | -34.3176 | -13.716
8.0027 | -34.3179 | -13.7259
8.0088 | -34.3182 | -13.7352
8.0149 | -34.3185 | -13.7445
8.021 | -34.3188 | -13.754
8.0271 | -34.3191 | -13.7638
8.0332 | -34.3193 | -13.7735
8.0393 | -34.3197 | -13.7831
8.0454 | -34.3199 | -13.7927
8.0515 | -34.3202 | -13.8022
8.0576 | -34.3205 | -13.8119
8.0637 | -34.3208 | -13.8211
8.0698 | -34.3211 | -13.8304
8.0759 | -34.3214 | -13.8397
8.082 | -34.3217 | -13.8496
8.0881 | -34.322 | -13.859
8.0942 | -34.3223 | -13.8689
8.1004 | -34.3226 | -13.8787
8.1065 | -34.3229 | -13.8884
8.1126 | -34.3233 | -13.8978
8.1187 | -34.3236 | -13.9074
8.1248 | -34.3239 | -13.9174
8.1309 | -34.3241 | -13.927
8.137 | -34.3244 | -13.9366
8.1431 | -34.3248 | -13.9462
8.1492 | -34.3251 | -13.9559
8.1553 | -34.3254 | -13.9654
8.1614 | -34.3257 | -13.975
8.1675 | -34.326 | -13.9845
8.1736 | -34.3263 | -13.9943
8.1797 | -34.3265 | -14.004
8.1858 | -34.3268 | -14.0135
8.1919 | -34.3271 | -14.023
8.198 | -34.3274 | -14.0323
8.2041 | -34.3277 | -14.0421
8.2102 | -34.3281 | -14.0517
8.2163 | -34.3284 | -14.0613
8.2224 | -34.3287 | -14.0711
8.2285 | -34.329 | -14.0806
8.2346 | -34.3293 | -14.0901
8.2408 | -34.3296 | -14.0999
8.2469 | -34.3299 | -14.1094
8.253 | -34.3301 | -14.1189
8.2591 | -34.3305 | -14.1281
8.2652 | -34.3308 | -14.1374
8.2713 | -34.3312 | -14.147
8.2774 | -34.3315 | -14.1566
8.2835 | -34.3318 | -14.1664
8.2896 | -34.3321 | -14.1756
8.2957 | -34.3324 | -14.1851
8.3018 | -34.3327 | -14.1945
8.3079 | -34.333 | -14.2041
8.314 | -34.3333 | -14.2135
8.3201 | -34.3336 | -14.2229
8.3262 | -34.3339 | -14.2324
8.3323 | -34.3343 | -14.2419
8.3384 | -34.3346 | -14.2518
8.3445 | -34.3349 | -14.261
8.3506 | -34.3352 | -14.2707
8.3567 | -34.3355 | -14.2804
8.3628 | -34.3358 | -14.2899
8.3689 | -34.3361 | -14.2995
8.375 | -34.3364 | -14.309
8.3812 | -34.3367 | -14.3187
8.3873 | -34.337 | -14.3283
8.3934 | -34.3373 | -14.3377
8.3995 | -34.3377 | -14.3475
8.4056 | -34.338 | -14.3571
8.4117 | -34.3383 | -14.3669
8.4178 | -34.3386 | -14.3761
8.4239 | -34.3389 | -14.3855
8.43 | -34.3393 | -14.395
8.4361 | -34.3396 | -14.4049
8.4422 | -34.3399 | -14.4144
8.4483 | -34.3402 | -14.424
8.4544 | -34.3405 | -14.4338
8.4605 | -34.3408 | -14.4433
8.4666 | -34.3411 | -14.4529
8.4727 | -34.3414 | -14.4624
8.4788 | -34.3417 | -14.4718
8.4849 | -34.342 | -14.4817
8.491 | -34.3423 | -14.491
8.4971 | -34.3426 | -14.5008
8.5032 | -34.3429 | -14.5105
8.5093 | -34.3432 | -14.52
8.5154 | -34.3435 | -14.5293
8.5215 | -34.3438 | -14.539
8.5277 | -34.3441 | -14.5486
8.5338 | -34.3444 | -14.5579
8.5399 | -34.3448 | -14.5678
8.546 | -34.3451 | -14.5771
8.5521 | -34.3454 | -14.5868
8.5582 | -34.3457 | -14.5965
8.5643 | -34.346 | -14.606
8.5704 | -34.3463 | -14.6154
8.5765 | -34.3466 | -14.6248
8.5826 | -34.347 | -14.6345
8.5887 | -34.3473 | -14.644
8.5948 | -34.3476 | -14.653
8.6009 | -34.3479 | -14.6627
8.607 | -34.3482 | -14.6724
8.6131 | -34.3484 | -14.6821
8.6192 | -34.3487 | -14.6913
8.6253 | -34.3491 | -14.701
8.6314 | -34.3494 | -14.7106
8.6375 | -34.3497 | -14.7203
8.6436 | -34.3501 | -14.7299
8.6497 | -34.3505 | -14.7395
8.6558 | -34.3508 | -14.7488
8.6619 | -34.3511 | -14.7585
8.6681 | -34.3513 | -14.7678
8.6742 | -34.3517 | -14.7776
8.6803 | -34.352 | -14.7874
8.6864 | -34.3524 | -14.7974
8.6925 | -34.3527 | -14.8071
8.6986 | -34.353 | -14.8169
8.7047 | -34.3534 | -14.8265
8.7108 | -34.3537 | -14.8361
8.7169 | -34.354 | -14.8455
8.723 | -34.3544 | -14.855
8.7291 | -34.3546 | -14.8647
8.7352 | -34.355 | -14.8744
8.7413 | -34.3553 | -14.8838
8.7474 | -34.3557 | -14.8932
8.7535 | -34.356 | -14.9029
8.7596 | -34.3564 | -14.9129
8.7657 | -34.3567 | -14.9224
8.7718 | -34.357 | -14.9317
8.7779 | -34.3574 | -14.9417
8.784 | -34.3576 | -14.9511
8.7901 | -34.3579 | -14.961
8.7962 | -34.3583 | -14.9708
8.8023 | -34.3587 | -14.9801
8.8084 | -34.359 | -14.9899
8.8146 | -34.3593 | -14.9994
8.8207 | -34.3596 | -15.009
8.8268 | -34.36 | -15.0186
8.8329 | -34.3603 | -15.0283
8.839 | -34.3606 | -15.0379
8.8451 | -34.361 | -15.0473
8.8512 | -34.3613 | -15.0571
8.8573 | -34.3617 | -15.0664
8.8634 | -34.3621 | -15.076
8.8695 | -34.3624 | -15.0856
8.8756 | -34.3627 | -15.095
8.8817 | -34.363 | -15.1048
8.8878 | -34.3633 | -15.1145
8.8939 | -34.3636 | -15.1238
8.9 | -34.364 | -15.133
8.9061 | -34.3643 | -15.1427
8.9122 | -34.3647 | -15.1521
8.9183 | -34.365 | -15.1615
8.9244 | -34.3653 | -15.1707
8.9305 | -34.3657 | -15.1804
8.9366 | -34.366 | -15.1898
8.9427 | -34.3664 | -15.1993
8.9488 | -34.3667 | -15.2089
8.955 | -34.3671 | -15.2185
8.9611 | -34.3674 | -15.228
8.9672 | -34.3678 | -15.2379
8.9733 | -34.3681 | -15.2474
8.9794 | -34.3684 | -15.2571
8.9855 | -34.3688 | -15.2665
8.9916 | -34.3691 | -15.2758
8.9977 | -34.3695 | -15.2853
9.0038 | -34.3698 | -15.295
9.0099 | -34.3702 | -15.3047
9.016 | -34.3706 | -15.3142
9.0221 | -34.3709 | -15.3242
9.0282 | -34.3713 | -15.334
9.0343 | -34.3716 | -15.3438
9.0404 | -34.372 | -15.3531
9.0465 | -34.3724 | -15.3629
9.0526 | -34.3727 | -15.3724
9.0587 | -34.3731 | -15.3818
9.0648 | -34.3735 | -15.3914
9.0709 | -34.3738 | -15.4007
9.077 | -34.3742 | -15.4107
9.0831 | -34.3745 | -15.4201
9.0892 | -34.3748 | -15.4297
9.0953 | -34.3751 | -15.4392
9.1015 | -34.3755 | -15.4489
9.1076 | -34.3759 | -15.4584
9.1137 | -34.3763 | -15.4679
9.1198 | -34.3766 | -15.4771
9.1259 | -34.377 | -15.4868
9.132 | -34.3774 | -15.4964
9.1381 | -34.3777 | -15.5059
9.1442 | -34.3781 | -15.5151
9.1503 | -34.3785 | -15.5247
9.1564 | -34.3788 | -15.5343
9.1625 | -34.3792 | -15.5438
9.1686 | -34.3795 | -15.5534
9.1747 | -34.3799 | -15.5631
9.1808 | -34.3802 | -15.5727
9.1869 | -34.3806 | -15.5826
9.193 | -34.3809 | -15.5923
9.1991 | -34.3813 | -15.6016
9.2052 | -34.3816 | -15.611
9.2113 | -34.382 | -15.6207
9.2174 | -34.3824 | -15.6307
9.2235 | -34.3828 | -15.64
9.2296 | -34.3831 | -15.6497
9.2357 | -34.3835 | -15.6591
9.2419 | -34.3839 | -15.6686
9.248 | -34.3842 | -15.6783
9.2541 | -34.3846 | -15.688
9.2602 | -34.3849 | -15.6975
9.2663 | -34.3853 | -15.7071
9.2724 | -34.3857 | -15.7167
9.2785 | -34.386 | -15.7263
9.2846 | -34.3864 | -15.7359
9.2907 | -34.3867 | -15.7453
9.2968 | -34.3871 | -15.7549
9.3029 | -34.3874 | -15.7644
9.309 | -34.3878 | -15.7739
9.3151 | -34.3882 | -15.7835
9.3212 | -34.3886 | -15.7931
9.3273 | -34.389 | -15.8026
9.3334 | -34.3894 | -15.812
9.3395 | -34.3898 | -15.8216
9.3456 | -34.3902 | -15.831
9.3517 | -34.3906 | -15.8401
9.3578 | -34.391 | -15.8494
9.3639 | -34.3914 | -15.8593
9.37 | -34.3918 | -15.8688
9.3761 | -34.3921 | -15.8782
9.3822 | -34.3925 | -15.8875
9.3884 | -34.3929 | -15.8972
9.3945 | -34.3933 | -15.9067
9.4006 | -34.3936 | -15.9163
9.4067 | -34.394 | -15.9259
9.4128 | -34.3944 | -15.9353
9.4189 | -34.3948 | -15.9448
9.425 | -34.3952 | -15.9542
9.4311 | -34.3955 | -15.9637
9.4372 | -34.3959 | -15.9732
9.4433 | -34.3963 | -15.9829
9.4494 | -34.3966 | -15.9925
9.4555 | -34.397 | -16.0019
9.4616 | -34.3974 | -16.0113
9.4677 | -34.3978 | -16.0207
9.4738 | -34.3982 | -16.0306
9.4799 | -34.3985 | -16.0399
9.486 | -34.3988 | -16.0498
9.4921 | -34.3993 | -16.0598
9.4982 | -34.3997 | -16.0688
9.5043 | -34.4 | -16.078
9.5104 | -34.4004 | -16.0875
9.5165 | -34.4008 | -16.097
9.5226 | -34.4012 | -16.1066
9.5288 | -34.4016 | -16.1161
9.5349 | -34.402 | -16.1255
9.541 | -34.4024 | -16.1351
9.5471 | -34.4028 | -16.1447
9.5532 | -34.4032 | -16.1541
9.5593 | -34.4036 | -16.1635
9.5654 | -34.404 | -16.1731
9.5715 | -34.4044 | -16.1828
9.5776 | -34.4048 | -16.1925
9.5837 | -34.4051 | -16.2018
9.5898 | -34.4055 | -16.2115
9.5959 | -34.4058 | -16.2208
9.602 | -34.4062 | -16.2305
9.6081 | -34.4066 | -16.2401
9.6142 | -34.407 | -16.2495
9.6203 | -34.4074 | -16.2591
9.6264 | -34.4077 | -16.2683
9.6325 | -34.408 | -16.278
9.6386 | -34.4085 | -16.2875
9.6447 | -34.4088 | -16.2968
9.6508 | -34.4092 | -16.3064
9.6569 | -34.4096 | -16.3159
9.663 | -34.41 | -16.3254
9.6691 | -34.4104 | -16.3349
9.6753 | -34.4109 | -16.3444
9.6814 | -34.4113 | -16.3538
9.6875 | -34.4117 | -16.3632
9.6936 | -34.4121 | -16.3727
9.6997 | -34.4125 | -16.3823
9.7058 | -34.4129 | -16.3917
9.7119 | -34.4132 | -16.401
9.718 | -34.4136 | -16.4107
9.7241 | -34.414 | -16.4202
9.7302 | -34.4143 | -16.4298
9.7363 | -34.4147 | -16.4391
9.7424 | -34.4152 | -16.4485
9.7485 | -34.4156 | -16.4579
9.7546 | -34.416 | -16.4673
9.7607 | -34.4163 | -16.4768
9.7668 | -34.4167 | -16.4867
9.7729 | -34.4171 | -16.496
9.779 | -34.4176 | -16.5055
9.7851 | -34.4179 | -16.515
9.7912 | -34.4183 | -16.525
9.7973 | -34.4186 | -16.5346
9.8034 | -34.4191 | -16.5444
9.8095 | -34.4194 | -16.5536
9.8157 | -34.4198 | -16.5631
9.8218 | -34.4202 | -16.5721
9.8279 | -34.4206 | -16.5818
9.834 | -34.4211 | -16.5911
9.8401 | -34.4214 | -16.6004
9.8462 | -34.4219 | -16.6097
9.8523 | -34.4223 | -16.6191
9.8584 | -34.4227 | -16.6288
9.8645 | -34.423 | -16.6383
9.8706 | -34.4234 | -16.6476
9.8767 | -34.4238 | -16.6567
9.8828 | -34.4243 | -16.6659
9.8889 | -34.4247 | -16.6757
9.895 | -34.425 | -16.6848
9.9011 | -34.4254 | -16.6938
9.9072 | -34.4258 | -16.7029
9.9133 | -34.4262 | -16.7122
9.9194 | -34.4267 | -16.7214
9.9255 | -34.427 | -16.7311
9.9316 | -34.4275 | -16.7402
9.9377 | -34.4279 | -16.7495
9.9438 | -34.4283 | -16.7588
9.9499 | -34.4287 | -16.7682
9.956 | -34.4291 | -16.7774
9.9622 | -34.4295 | -16.7866
9.9683 | -34.4299 | -16.7961
9.9744 | -34.4303 | -16.8053
9.9805 | -34.4307 | -16.8147
9.9866 | -34.4311 | -16.8239
9.9927 | -34.4316 | -16.8336
9.9988 | -34.432 | -16.8427
10.0049 | -34.4323 | -16.852
10.011 | -34.4328 | -16.8613
10.0171 | -34.4331 | -16.8707
10.0232 | -34.4336 | -16.8799
10.0293 | -34.434 | -16.889
10.0354 | -34.4345 | -16.898
10.0415 | -34.4348 | -16.9073
10.0476 | -34.4352 | -16.9169
10.0537 | -34.4356 | -16.9264
10.0598 | -34.4361 | -16.936
10.0659 | -34.4364 | -16.9453
10.072 | -34.4369 | -16.9546
10.0781 | -34.4373 | -16.964
10.0842 | -34.4377 | -16.9735
10.0903 | -34.4381 | -16.9831
10.0964 | -34.4386 | -16.9923
10.1026 | -34.439 | -17.0016
10.1087 | -34.4395 | -17.011
10.1148 | -34.4399 | -17.0204
10.1209 | -34.4403 | -17.0297
10.127 | -34.4407 | -17.0389
10.1331 | -34.4411 | -17.0482
10.1392 | -34.4415 | -17.0576
10.1453 | -34.4419 | -17.0669
10.1514 | -34.4423 | -17.076
10.1575 | -34.4427 | -17.0851
10.1636 | -34.4432 | -17.0939
10.1697 | -34.4436 | -17.1032
10.1758 | -34.444 | -17.1124
10.1819 | -34.4444 | -17.1219
10.188 | -34.4449 | -17.1309
10.1941 | -34.4452 | -17.1403
10.2002 | -34.4456 | -17.149
10.2063 | -34.446 | -17.1586
10.2124 | -34.4464 | -17.1678
10.2185 | -34.4468 | -17.177
10.2246 | -34.4472 | -17.1863
10.2307 | -34.4476 | -17.1958
10.2368 | -34.4481 | -17.2044
10.2429 | -34.4485 | -17.2134
10.2491 | -34.4489 | -17.2224
10.2552 | -34.4493 | -17.2321
10.2613 | -34.4498 | -17.2416
10.2674 | -34.4501 | -17.2509
10.2735 | -34.4505 | -17.2603
10.2796 | -34.4509 | -17.2697
10.2857 | -34.4513 | -17.2789
10.2918 | -34.4517 | -17.2881
10.2979 | -34.4522 | -17.2976
10.304 | -34.4526 | -17.3071
10.3101 | -34.4531 | -17.3163
10.3162 | -34.4535 | -17.3255
10.3223 | -34.4539 | -17.3348
10.3284 | -34.4543 | -17.344
10.3345 | -34.4547 | -17.3532
10.3406 | -34.4551 | -17.3623
10.3467 | -34.4555 | -17.3716
10.3528 | -34.4559 | -17.3807
10.3589 | -34.4563 | -17.3901
10.365 | -34.4567 | -17.3995
10.3711 | -34.4571 | -17.4089
10.3772 | -34.4574 | -17.4182
10.3833 | -34.4578 | -17.4274
10.3895 | -34.4582 | -17.4364
10.3956 | -34.4587 | -17.4454
10.4017 | -34.459 | -17.4547
10.4078 | -34.4595 | -17.464
10.4139 | -34.4599 | -17.4731
10.42 | -34.4603 | -17.4823
10.4261 | -34.4608 | -17.4915
10.4322 | -34.4612 | -17.5006
10.4383 | -34.4616 | -17.5099
10.4444 | -34.462 | -17.5188
10.4505 | -34.4625 | -17.528
10.4566 | -34.4629 | -17.5372
10.4627 | -34.4633 | -17.5465
10.4688 | -34.4637 | -17.5553
10.4749 | -34.4641 | -17.5643
10.481 | -34.4645 | -17.5736
10.4871 | -34.4649 | -17.583
10.4932 | -34.4654 | -17.5921
10.4993 | -34.4658 | -17.6012
10.5054 | -34.4662 | -17.6105
10.5115 | -34.4666 | -17.6197
10.5176 | -34.4671 | -17.6286
10.5237 | -34.4675 | -17.6376
10.5298 | -34.4679 | -17.6467
10.536 | -34.4684 | -17.656
10.5421 | -34.4687 | -17.6648
10.5482 | -34.4692 | -17.6743
10.5543 | -34.4696 | -17.6834
10.5604 | -34.47 | -17.6923
10.5665 | -34.4705 | -17.7014
10.5726 | -34.4708 | -17.7104
10.5787 | -34.4713 | -17.7197
10.5848 | -34.4718 | -17.7287
10.5909 | -34.4721 | -17.7379
10.597 | -34.4725 | -17.7469
10.6031 | -34.473 | -17.7561
10.6092 | -34.4734 | -17.7656
10.6153 | -34.4738 | -17.7749
10.6214 | -34.4742 | -17.784
10.6275 | -34.4747 | -17.7931
10.6336 | -34.4751 | -17.8025
10.6397 | -34.4755 | -17.8116
10.6458 | -34.4759 | -17.8205
10.6519 | -34.4763 | -17.8295
10.658 | -34.4767 | -17.8383
10.6641 | -34.4771 | -17.8474
10.6702 | -34.4775 | -17.8567
10.6764 | -34.478 | -17.8657
10.6825 | -34.4784 | -17.8749
10.6886 | -34.4789 | -17.8843
10.6947 | -34.4793 | -17.8931
10.7008 | -34.4797 | -17.9022
10.7069 | -34.4801 | -17.9113
10.713 | -34.4806 | -17.9203
10.7191 | -34.481 | -17.9296
10.7252 | -34.4815 | -17.9386
10.7313 | -34.4819 | -17.9478
10.7374 | -34.4823 | -17.9566
10.7435 | -34.4827 | -17.966
10.7496 | -34.4831 | -17.975
10.7557 | -34.4836 | -17.9841
10.7618 | -34.484 | -17.9932
10.7679 | -34.4844 | -18.0023
10.774 | -34.4848 | -18.0113
10.7801 | -34.4852 | -18.0209
10.7862 | -34.4857 | -18.03
10.7923 | -34.4861 | -18.0391
10.7984 | -34.4865 | -18.048
10.8045 | -34.4869 | -18.0574
10.8106 | -34.4874 | -18.0664
10.8168 | -34.4878 | -18.0756
10.8229 | -34.4882 | -18.0848
10.829 | -34.4887 | -18.094
10.8351 | -34.4891 | -18.1032
10.8412 | -34.4894 | -18.1122
10.8473 | -34.4898 | -18.1212
10.8534 | -34.4902 | -18.1304
10.8595 | -34.4907 | -18.1393
10.8656 | -34.4911 | -18.1484
10.8717 | -34.4916 | -18.1572
10.8778 | -34.492 | -18.1661
10.8839 | -34.4924 | -18.1751
10.89 | -34.4928 | -18.184
10.8961 | -34.4932 | -18.1931
10.9022 | -34.4937 | -18.2021
10.9083 | -34.4941 | -18.2115
10.9144 | -34.4946 | -18.2205
10.9205 | -34.495 | -18.2295
10.9266 | -34.4955 | -18.2388
10.9327 | -34.4959 | -18.2479
10.9388 | -34.4963 | -18.2569
10.9449 | -34.4967 | -18.266
10.951 | -34.4971 | -18.275
10.9571 | -34.4976 | -18.2841
10.9633 | -34.498 | -18.2934
10.9694 | -34.4985 | -18.3027
10.9755 | -34.4989 | -18.3117
10.9816 | -34.4993 | -18.3207
10.9877 | -34.4998 | -18.3299
10.9938 | -34.5002 | -18.339
10.9999 | -34.5006 | -18.348
11.006 | -34.5011 | -18.357
11.0121 | -34.5015 | -18.3665
11.0182 | -34.5019 | -18.3751
11.0243 | -34.5024 | -18.3842
11.0304 | -34.5027 | -18.3932
11.0365 | -34.5032 | -18.4022
11.0426 | -34.5036 | -18.4111
11.0487 | -34.504 | -18.4201
11.0548 | -34.5045 | -18.4292
11.0609 | -34.5049 | -18.4385
11.067 | -34.5053 | -18.4478
11.0731 | -34.5058 | -18.4569
11.0792 | -34.5062 | -18.4659
11.0853 | -34.5066 | -18.4746
11.0914 | -34.507 | -18.4835
11.0975 | -34.5075 | -18.4929
11.1037 | -34.5079 | -18.5016
11.1098 | -34.5083 | -18.5108
11.1159 | -34.5088 | -18.52
11.122 | -34.5092 | -18.5294
11.1281 | -34.5096 | -18.5382
11.1342 | -34.51 | -18.5469
11.1403 | -34.5104 | -18.5562
11.1464 | -34.5109 | -18.5651
11.1525 | -34.5113 | -18.5745
11.1586 | -34.5118 | -18.584
11.1647 | -34.5122 | -18.5929
11.1708 | -34.5126 | -18.6018
11.1769 | -34.513 | -18.6104
11.183 | -34.5135 | -18.6193
11.1891 | -34.5139 | -18.6287
11.1952 | -34.5143 | -18.638
11.2013 | -34.5147 | -18.647
11.2074 | -34.5151 | -18.6558
11.2135 | -34.5155 | -18.6648
11.2196 | -34.5159 | -18.6738
11.2257 | -34.5163 | -18.6826
11.2318 | -34.5166 | -18.6919
11.2379 | -34.517 | -18.7007
11.244 | -34.5174 | -18.7096
11.2502 | -34.5178 | -18.7185
11.2563 | -34.5183 | -18.7274
11.2624 | -34.5186 | -18.7366
11.2685 | -34.5191 | -18.7454
11.2746 | -34.5195 | -18.7542
11.2807 | -34.5199 | -18.7635
11.2868 | -34.5202 | -18.7727
11.2929 | -34.5207 | -18.7816
11.299 | -34.5211 | -18.7907
11.3051 | -34.5215 | -18.7996
11.3112 | -34.5219 | -18.8086
11.3173 | -34.5224 | -18.8178
11.3234 | -34.5227 | -18.8269
11.3295 | -34.5232 | -18.8356
11.3356 | -34.5236 | -18.8444
11.3417 | -34.524 | -18.8536
11.3478 | -34.5244 | -18.8626
11.3539 | -34.5248 | -18.8715
11.36 | -34.5252 | -18.8804
11.3661 | -34.5257 | -18.8893
11.3722 | -34.5261 | -18.8983
11.3783 | -34.5265 | -18.9072
11.3844 | -34.5269 | -18.9162
11.3906 | -34.5273 | -18.9255
11.3967 | -34.5277 | -18.9343
11.4028 | -34.5281 | -18.9434
11.4089 | -34.5285 | -18.9524
11.415 | -34.5289 | -18.9613
11.4211 | -34.5294 | -18.97
11.4272 | -34.5298 | -18.979
11.4333 | -34.5302 | -18.9881
11.4394 | -34.5306 | -18.9974
11.4455 | -34.531 | -19.0065
11.4516 | -34.5314 | -19.0153
11.4577 | -34.5319 | -19.0245
11.4638 | -34.5323 | -19.0331
11.4699 | -34.5327 | -19.0424
11.476 | -34.5331 | -19.0517
11.4821 | -34.5335 | -19.0604
11.4882 | -34.534 | -19.0696
11.4943 | -34.5344 | -19.0784
11.5004 | -34.5348 | -19.0874
11.5065 | -34.5353 | -19.0964
11.5126 | -34.5357 | -19.1052
11.5187 | -34.5361 | -19.1141
11.5248 | -34.5365 | -19.123
11.5309 | -34.5369 | -19.132
11.5371 | -34.5373 | -19.141
11.5432 | -34.5377 | -19.1501
11.5493 | -34.5381 | -19.1591
11.5554 | -34.5385 | -19.1679
11.5615 | -34.5389 | -19.177
11.5676 | -34.5394 | -19.1861
11.5737 | -34.5398 | -19.1949
11.5798 | -34.5402 | -19.2037
11.5859 | -34.5407 | -19.2126
11.592 | -34.5411 | -19.2216
11.5981 | -34.5415 | -19.2305
11.6042 | -34.5419 | -19.2395
11.6103 | -34.5423 | -19.2485
11.6164 | -34.5427 | -19.2573
11.6225 | -34.5431 | -19.2662
11.6286 | -34.5436 | -19.2753
11.6347 | -34.544 | -19.2844
11.6408 | -34.5444 | -19.2934
11.6469 | -34.5449 | -19.3025
11.653 | -34.5453 | -19.3113
11.6591 | -34.5458 | -19.3203
11.6652 | -34.5462 | -19.3295
11.6713 | -34.5465 | -19.3384
11.6775 | -34.547 | -19.3473
11.6836 | -34.5474 | -19.3563
11.6897 | -34.5478 | -19.3651
11.6958 | -34.5483 | -19.3743
11.7019 | -34.5487 | -19.3832
11.708 | -34.5491 | -19.3924
11.7141 | -34.5496 | -19.4014
11.7202 | -34.55 | -19.4101
11.7263 | -34.5504 | -19.4191
11.7324 | -34.5509 | -19.4281
11.7385 | -34.5513 | -19.4372
11.7446 | -34.5517 | -19.4464
11.7507 | -34.5521 | -19.4553
11.7568 | -34.5525 | -19.4643
11.7629 | -34.553 | -19.4733
11.769 | -34.5534 | -19.4822
11.7751 | -34.5538 | -19.4909
11.7812 | -34.5542 | -19.4997
11.7873 | -34.5547 | -19.5089
11.7934 | -34.5551 | -19.5178
11.7995 | -34.5555 | -19.5265
11.8056 | -34.5559 | -19.5359
11.8117 | -34.5563 | -19.5449
11.8178 | -34.5567 | -19.5537
11.824 | -34.5572 | -19.5626
11.8301 | -34.5577 | -19.5712
11.8362 | -34.5581 | -19.5799
11.8423 | -34.5585 | -19.5889
11.8484 | -34.5589 | -19.5978
11.8545 | -34.5593 | -19.607
11.8606 | -34.5597 | -19.6161
11.8667 | -34.5602 | -19.625
11.8728 | -34.5606 | -19.634
11.8789 | -34.561 | -19.6434
11.885 | -34.5613 | -19.6522
11.8911 | -34.5618 | -19.6613
11.8972 | -34.5622 | -19.67
11.9033 | -34.5626 | -19.6785
11.9094 | -34.563 | -19.6872
11.9155 | -34.5634 | -19.6962
11.9216 | -34.5639 | -19.705
11.9277 | -34.5643 | -19.7138
11.9338 | -34.5647 | -19.7227
11.9399 | -34.5651 | -19.7318
11.946 | -34.5656 | -19.7408
11.9521 | -34.566 | -19.7496
11.9582 | -34.5664 | -19.7583
11.9644 | -34.5669 | -19.7671
11.9705 | -34.5673 | -19.7761
11.9766 | -34.5677 | -19.785
11.9827 | -34.5681 | -19.7938
11.9888 | -34.5685 | -19.8025
11.9949 | -34.5689 | -19.8115
12.001 | -34.5694 | -19.8204
12.0071 | -34.5698 | -19.8291
12.0132 | -34.5702 | -19.8377
12.0193 | -34.5707 | -19.8468
12.0254 | -34.5711 | -19.8555
12.0315 | -34.5715 | -19.8644
12.0376 | -34.5719 | -19.8731
12.0437 | -34.5723 | -19.8819
12.0498 | -34.5728 | -19.8909
12.0559 | -34.5732 | -19.8999
12.062 | -34.5736 | -19.9089
12.0681 | -34.574 | -19.9177
12.0742 | -34.5744 | -19.9268
12.0803 | -34.5749 | -19.9355
12.0864 | -34.5753 | -19.9443
12.0925 | -34.5757 | -19.9532
12.0986 | -34.5762 | -19.9618
12.1047 | -34.5767 | -19.9706
12.1109 | -34.5771 | -19.9797
12.117 | -34.5775 | -19.9885
12.1231 | -34.578 | -19.9974
12.1292 | -34.5784 | -20.0065
12.1353 | -34.5789 | -20.0151
12.1414 | -34.5793 | -20.024
12.1475 | -34.5797 | -20.0329
12.1536 | -34.5801 | -20.0418
12.1597 | -34.5805 | -20.0508
12.1658 | -34.5809 | -20.0597
12.1719 | -34.5814 | -20.0688
12.178 | -34.5818 | -20.0778
12.1841 | -34.5822 | -20.0868
12.1902 | -34.5827 | -20.0957
12.1963 | -34.5831 | -20.1045
12.2024 | -34.5835 | -20.1132
12.2085 | -34.5839 | -20.1222
12.2146 | -34.5844 | -20.1308
12.2207 | -34.5848 | -20.1398
12.2268 | -34.5852 | -20.1486
12.2329 | -34.5856 | -20.1574
12.239 | -34.586 | -20.1661
12.2451 | -34.5865 | -20.1747
12.2513 | -34.5869 | -20.1836
12.2574 | -34.5874 | -20.1926
12.2635 | -34.5878 | -20.2011
12.2696 | -34.5882 | -20.2099
12.2757 | -34.5886 | -20.2184
12.2818 | -34.589 | -20.227
12.2879 | -34.5894 | -20.2356
12.294 | -34.5898 | -20.2444
12.3001 | -34.5903 | -20.2535
12.3062 | -34.5907 | -20.2622
12.3123 | -34.5911 | -20.2711
12.3184 | -34.5916 | -20.2798
12.3245 | -34.592 | -20.2888
12.3306 | -34.5924 | -20.2974
12.3367 | -34.5928 | -20.3061
12.3428 | -34.5933 | -20.3146
12.3489 | -34.5937 | -20.323
12.355 | -34.5941 | -20.3317
12.3611 | -34.5945 | -20.3404
12.3672 | -34.595 | -20.3491
12.3733 | -34.5954 | -20.3581
12.3794 | -34.5958 | -20.3665
12.3855 | -34.5962 | -20.3753
12.3916 | -34.5966 | -20.3842
12.3978 | -34.5971 | -20.3932
12.4039 | -34.5974 | -20.402
12.41 | -34.5979 | -20.4108
12.4161 | -34.5983 | -20.4195
12.4222 | -34.5988 | -20.4281
12.4283 | -34.5992 | -20.437
12.4344 | -34.5996 | -20.4457
12.4405 | -34.6 | -20.4544
12.4466 | -34.6004 | -20.4631
12.4527 | -34.6009 | -20.4719
12.4588 | -34.6013 | -20.4806
12.4649 | -34.6017 | -20.4896
12.471 | -34.6022 | -20.4988
12.4771 | -34.6026 | -20.507
12.4832 | -34.603 | -20.5158
12.4893 | -34.6035 | -20.5243
12.4954 | -34.6039 | -20.533
12.5015 | -34.6044 | -20.5416
12.5076 | -34.6048 | -20.5503
12.5137 | -34.6052 | -20.5594
12.5198 | -34.6056 | -20.5679
12.5259 | -34.6061 | -20.5764
12.532 | -34.6066 | -20.5854
12.5382 | -34.607 | -20.5939
12.5443 | -34.6074 | -20.6028
12.5504 | -34.6078 | -20.6116
12.5565 | -34.6083 | -20.62
12.5626 | -34.6088 | -20.6289
12.5687 | -34.6092 | -20.6378
12.5748 | -34.6096 | -20.6464
12.5809 | -34.61 | -20.6548
12.587 | -34.6104 | -20.6638
12.5931 | -34.6109 | -20.6725
12.5992 | -34.6113 | -20.6814
12.6053 | -34.6117 | -20.69
12.6114 | -34.6122 | -20.6988
12.6175 | -34.6126 | -20.7075
12.6236 | -34.613 | -20.716
12.6297 | -34.6135 | -20.7246
12.6358 | -34.6139 | -20.7331
12.6419 | -34.6144 | -20.742
12.648 | -34.6148 | -20.7509
12.6541 | -34.6153 | -20.7595
12.6602 | -34.6157 | -20.7678
12.6663 | -34.6161 | -20.7763
12.6724 | -34.6166 | -20.7851
12.6785 | -34.617 | -20.794
12.6847 | -34.6174 | -20.8028
12.6908 | -34.6179 | -20.8114
12.6969 | -34.6183 | -20.82
12.703 | -34.6188 | -20.8287
12.7091 | -34.6192 | -20.8375
12.7152 | -34.6196 | -20.8462
12.7213 | -34.6201 | -20.8551
12.7274 | -34.6205 | -20.8637
12.7335 | -34.6209 | -20.872
12.7396 | -34.6213 | -20.8804
12.7457 | -34.6218 | -20.8892
12.7518 | -34.6222 | -20.8979
12.7579 | -34.6226 | -20.9068
12.764 | -34.623 | -20.9155
12.7701 | -34.6234 | -20.9237
12.7762 | -34.6238 | -20.9326
12.7823 | -34.6242 | -20.9411
12.7884 | -34.6246 | -20.9496
12.7945 | -34.6251 | -20.9578
12.8006 | -34.6255 | -20.9659
12.8067 | -34.6259 | -20.9748
12.8128 | -34.6263 | -20.9831
12.8189 | -34.6267 | -20.9915
12.8251 | -34.627 | -21
12.8312 | -34.6275 | -21.0086
12.8373 | -34.6279 | -21.0171
12.8434 | -34.6284 | -21.0256
12.8495 | -34.6288 | -21.0342
12.8556 | -34.6292 | -21.0429
12.8617 | -34.6296 | -21.0512
12.8678 | -34.63 | -21.0597
12.8739 | -34.6305 | -21.068
12.88 | -34.6309 | -21.0767
12.8861 | -34.6313 | -21.0853
12.8922 | -34.6317 | -21.0938
12.8983 | -34.6321 | -21.102
12.9044 | -34.6326 | -21.1107
12.9105 | -34.6331 | -21.1191
12.9166 | -34.6335 | -21.1281
12.9227 | -34.6339 | -21.1371
12.9288 | -34.6344 | -21.1458
12.9349 | -34.6349 | -21.1548
12.941 | -34.6353 | -21.1634
12.9471 | -34.6358 | -21.1718
12.9532 | -34.6363 | -21.1804
12.9593 | -34.6367 | -21.1891
12.9654 | -34.6371 | -21.1978
12.9716 | -34.6375 | -21.2063
12.9777 | -34.6379 | -21.2144
12.9838 | -34.6383 | -21.2229
12.9899 | -34.6388 | -21.2315
12.996 | -34.6391 | -21.2402
13.0021 | -34.6395 | -21.2488
13.0082 | -34.64 | -21.2574
13.0143 | -34.6403 | -21.2655
13.0204 | -34.6407 | -21.2738
13.0265 | -34.6412 | -21.2821
13.0326 | -34.6416 | -21.2909
13.0387 | -34.6421 | -21.2995
13.0448 | -34.6425 | -21.308
13.0509 | -34.6429 | -21.3164
13.057 | -34.6434 | -21.3246
13.0631 | -34.6439 | -21.333
13.0692 | -34.6443 | -21.3414
13.0753 | -34.6447 | -21.3498
13.0814 | -34.6451 | -21.3579
13.0875 | -34.6455 | -21.3661
13.0936 | -34.646 | -21.3744
13.0997 | -34.6464 | -21.3826
13.1058 | -34.6468 | -21.3904
13.112 | -34.6472 | -21.3988
13.1181 | -34.6476 | -21.4074
13.1242 | -34.648 | -21.4158
13.1303 | -34.6485 | -21.4244
13.1364 | -34.6489 | -21.433
13.1425 | -34.6493 | -21.4414
13.1486 | -34.6498 | -21.4497
13.1547 | -34.6502 | -21.4578
13.1608 | -34.6506 | -21.4662
13.1669 | -34.6511 | -21.4746
13.173 | -34.6515 | -21.483
13.1791 | -34.6519 | -21.4915
13.1852 | -34.6523 | -21.5003
13.1913 | -34.6527 | -21.5085
13.1974 | -34.6531 | -21.517
13.2035 | -34.6536 | -21.5254
13.2096 | -34.654 | -21.5339
13.2157 | -34.6544 | -21.5424
13.2218 | -34.6549 | -21.551
13.2279 | -34.6553 | -21.5593
13.234 | -34.6558 | -21.568
13.2401 | -34.6562 | -21.5763
13.2462 | -34.6566 | -21.5848
13.2524 | -34.657 | -21.593
13.2585 | -34.6574 | -21.6014
13.2646 | -34.6579 | -21.61
13.2707 | -34.6583 | -21.6185
13.2768 | -34.6587 | -21.6271
13.2829 | -34.6591 | -21.6356
13.289 | -34.6595 | -21.644
13.2951 | -34.6599 | -21.6523
13.3012 | -34.6604 | -21.6608
13.3073 | -34.6608 | -21.6694
13.3134 | -34.6612 | -21.6776
13.3195 | -34.6616 | -21.6859
13.3256 | -34.662 | -21.6944
13.3317 | -34.6625 | -21.7026
13.3378 | -34.6629 | -21.711
13.3439 | -34.6633 | -21.7192
13.35 | -34.6637 | -21.7275
13.3561 | -34.6641 | -21.7362
13.3622 | -34.6645 | -21.7446
13.3683 | -34.6649 | -21.7529
13.3744 | -34.6654 | -21.7616
13.3805 | -34.6658 | -21.7702
13.3866 | -34.6662 | -21.7785
13.3927 | -34.6666 | -21.7868
13.3989 | -34.667 | -21.7951
13.405 | -34.6675 | -21.8034
13.4111 | -34.6679 | -21.8118
13.4172 | -34.6683 | -21.82
13.4233 | -34.6687 | -21.8285
13.4294 | -34.6692 | -21.8366
13.4355 | -34.6696 | -21.8455
13.4416 | -34.67 | -21.8539
13.4477 | -34.6704 | -21.8622
13.4538 | -34.6708 | -21.8705
13.4599 | -34.6712 | -21.8789
13.466 | -34.6717 | -21.8877
13.4721 | -34.6721 | -21.8962
13.4782 | -34.6726 | -21.9045
13.4843 | -34.6731 | -21.9128
13.4904 | -34.6735 | -21.9207
13.4965 | -34.6739 | -21.9294
13.5026 | -34.6743 | -21.9378
13.5087 | -34.6748 | -21.9458
13.5148 | -34.6751 | -21.9542
13.5209 | -34.6756 | -21.9622
13.527 | -34.676 | -21.971
13.5331 | -34.6764 | -21.9793
13.5393 | -34.6768 | -21.9877
13.5454 | -34.6772 | -21.9962
13.5515 | -34.6777 | -22.0046
13.5576 | -34.6781 | -22.0131
13.5637 | -34.6785 | -22.0213
13.5698 | -34.679 | -22.03
13.5759 | -34.6793 | -22.0378
13.582 | -34.6797 | -22.0463
13.5881 | -34.6801 | -22.0547
13.5942 | -34.6805 | -22.0629
13.6003 | -34.681 | -22.0713
13.6064 | -34.6814 | -22.0799
13.6125 | -34.6818 | -22.0887
13.6186 | -34.6823 | -22.0972
13.6247 | -34.6827 | -22.1054
13.6308 | -34.6831 | -22.1133
13.6369 | -34.6836 | -22.1216
13.643 | -34.684 | -22.1298
13.6491 | -34.6845 | -22.1384
13.6552 | -34.6849 | -22.1465
13.6613 | -34.6853 | -22.1549
13.6674 | -34.6857 | -22.1634
13.6735 | -34.6861 | -22.1719
13.6796 | -34.6866 | -22.1799
13.6858 | -34.6869 | -22.1884
13.6919 | -34.6874 | -22.1967
13.698 | -34.6878 | -22.205
13.7041 | -34.6881 | -22.2135
13.7102 | -34.6886 | -22.222
13.7163 | -34.6891 | -22.2303
13.7224 | -34.6895 | -22.2384
13.7285 | -34.6899 | -22.2468
13.7346 | -34.6903 | -22.2551
13.7407 | -34.6907 | -22.2634
13.7468 | -34.6911 | -22.2717
13.7529 | -34.6915 | -22.2799
13.759 | -34.692 | -22.2883
13.7651 | -34.6924 | -22.2967
13.7712 | -34.6928 | -22.3053
13.7773 | -34.6932 | -22.3137
13.7834 | -34.6937 | -22.322
13.7895 | -34.6941 | -22.3301
13.7956 | -34.6946 | -22.3386
13.8017 | -34.695 | -22.3469
13.8078 | -34.6954 | -22.3551
13.8139 | -34.6959 | -22.3634
13.82 | -34.6963 | -22.3715
13.8262 | -34.6968 | -22.3801
13.8323 | -34.6972 | -22.3885
13.8384 | -34.6977 | -22.397
13.8445 | -34.6981 | -22.4055
13.8506 | -34.6985 | -22.4138
13.8567 | -34.6989 | -22.4224
13.8628 | -34.6993 | -22.4308
13.8689 | -34.6997 | -22.4392
13.875 | -34.7002 | -22.4474
13.8811 | -34.7007 | -22.4557
13.8872 | -34.7011 | -22.464
13.8933 | -34.7016 | -22.4721
13.8994 | -34.702 | -22.4801
13.9055 | -34.7024 | -22.4884
13.9116 | -34.7029 | -22.4965
13.9177 | -34.7034 | -22.5044
13.9238 | -34.7038 | -22.5128
13.9299 | -34.7043 | -22.5212
13.936 | -34.7047 | -22.5293
13.9421 | -34.7051 | -22.5377
13.9482 | -34.7055 | -22.5461
13.9543 | -34.706 | -22.5541
13.9604 | -34.7064 | -22.5626
13.9665 | -34.7068 | -22.5707
13.9727 | -34.7073 | -22.5789
13.9788 | -34.7077 | -22.5874
13.9849 | -34.708 | -22.5958
13.991 | -34.7084 | -22.6043
13.9971 | -34.7089 | -22.6129
14.0032 | -34.7093 | -22.6216
14.0093 | -34.7097 | -22.6297
14.0154 | -34.7101 | -22.6382
14.0215 | -34.7105 | -22.6463
14.0276 | -34.7109 | -22.6543
14.0337 | -34.7114 | -22.6627
14.0398 | -34.7118 | -22.6708
14.0459 | -34.7122 | -22.6791
14.052 | -34.7126 | -22.6874
14.0581 | -34.713 | -22.6954
14.0642 | -34.7135 | -22.704
14.0703 | -34.7139 | -22.7123
14.0764 | -34.7143 | -22.721
14.0825 | -34.7148 | -22.7294
14.0886 | -34.7152 | -22.7372
14.0947 | -34.7155 | -22.7451
14.1008 | -34.716 | -22.7538
14.1069 | -34.7164 | -22.762
14.1131 | -34.7168 | -22.7706
14.1192 | -34.7172 | -22.7787
14.1253 | -34.7176 | -22.7866
14.1314 | -34.7181 | -22.7949
14.1375 | -34.7185 | -22.8033
14.1436 | -34.719 | -22.8116
14.1497 | -34.7194 | -22.8197
14.1558 | -34.7198 | -22.828
14.1619 | -34.7203 | -22.8363
14.168 | -34.7207 | -22.8444
14.1741 | -34.7212 | -22.8525
14.1802 | -34.7216 | -22.8607
14.1863 | -34.722 | -22.8689
14.1924 | -34.7225 | -22.8771
14.1985 | -34.7229 | -22.8851
14.2046 | -34.7233 | -22.8932
14.2107 | -34.7238 | -22.9011
14.2168 | -34.7242 | -22.9097
14.2229 | -34.7247 | -22.918
14.229 | -34.7251 | -22.9261
14.2351 | -34.7255 | -22.9341
14.2412 | -34.726 | -22.9421
14.2473 | -34.7264 | -22.9504
14.2534 | -34.7269 | -22.9586
14.2596 | -34.7273 | -22.9669
14.2657 | -34.7277 | -22.9754
14.2718 | -34.7282 | -22.9837
14.2779 | -34.7286 | -22.9917
14.284 | -34.729 | -22.9997
14.2901 | -34.7295 | -23.0078
14.2962 | -34.73 | -23.0158
14.3023 | -34.7304 | -23.0241
14.3084 | -34.7309 | -23.0326
14.3145 | -34.7313 | -23.0406
14.3206 | -34.7317 | -23.0487
14.3267 | -34.7322 | -23.0568
14.3328 | -34.7326 | -23.0649
14.3389 | -34.733 | -23.073
14.345 | -34.7334 | -23.0813
14.3511 | -34.7339 | -23.089
14.3572 | -34.7343 | -23.0974
14.3633 | -34.7347 | -23.1055
14.3694 | -34.7351 | -23.1136
14.3755 | -34.7355 | -23.1214
14.3816 | -34.736 | -23.1294
14.3877 | -34.7365 | -23.1374
14.3938 | -34.7369 | -23.1457
14.4 | -34.7373 | -23.1536
14.4061 | -34.7377 | -23.1618
14.4122 | -34.7381 | -23.1702
14.4183 | -34.7385 | -23.1784
14.4244 | -34.7389 | -23.1864
14.4305 | -34.7394 | -23.1949
14.4366 | -34.7398 | -23.2032
14.4427 | -34.7403 | -23.2114
14.4488 | -34.7407 | -23.2192
14.4549 | -34.7412 | -23.2272
14.461 | -34.7416 | -23.2355
14.4671 | -34.742 | -23.2438
14.4732 | -34.7425 | -23.2517
14.4793 | -34.7429 | -23.2603
14.4854 | -34.7434 | -23.2681
14.4915 | -34.7438 | -23.2761
14.4976 | -34.7443 | -23.2843
14.5037 | -34.7448 | -23.2927
14.5098 | -34.7452 | -23.3006
14.5159 | -34.7457 | -23.309
14.522 | -34.7461 | -23.3173
14.5281 | -34.7466 | -23.325
14.5342 | -34.747 | -23.3332
14.5403 | -34.7474 | -23.3414
14.5465 | -34.7479 | -23.3499
14.5526 | -34.7484 | -23.3583
14.5587 | -34.7488 | -23.3663
14.5648 | -34.7492 | -23.374
14.5709 | -34.7497 | -23.3823
14.577 | -34.7501 | -23.3905
14.5831 | -34.7506 | -23.3987
14.5892 | -34.7511 | -23.4064
14.5953 | -34.7515 | -23.4139
14.6014 | -34.752 | -23.422
14.6075 | -34.7524 | -23.4298
14.6136 | -34.7529 | -23.4377
14.6197 | -34.7533 | -23.4459
14.6258 | -34.7538 | -23.454
14.6319 | -34.7543 | -23.4618
14.638 | -34.7547 | -23.4699
14.6441 | -34.7552 | -23.4776
14.6502 | -34.7556 | -23.4853
14.6563 | -34.756 | -23.4933
14.6624 | -34.7565 | -23.5013
14.6685 | -34.7569 | -23.5089
14.6746 | -34.7574 | -23.517
14.6807 | -34.7578 | -23.5251
14.6869 | -34.7582 | -23.533
14.693 | -34.7587 | -23.541
14.6991 | -34.7591 | -23.5488
14.7052 | -34.7596 | -23.5567
14.7113 | -34.76 | -23.5648
14.7174 | -34.7604 | -23.5728
14.7235 | -34.7609 | -23.5808
14.7296 | -34.7613 | -23.5887
14.7357 | -34.7618 | -23.5964
14.7418 | -34.7622 | -23.6041
14.7479 | -34.7627 | -23.6121
14.754 | -34.7632 | -23.6201
14.7601 | -34.7636 | -23.6284
14.7662 | -34.764 | -23.6364
14.7723 | -34.7645 | -23.6442
14.7784 | -34.7649 | -23.6523
14.7845 | -34.7654 | -23.6603
14.7906 | -34.7658 | -23.6683
14.7967 | -34.7663 | -23.6762
14.8028 | -34.7667 | -23.6841
14.8089 | -34.7671 | -23.6918
14.815 | -34.7676 | -23.6996
14.8211 | -34.768 | -23.7075
14.8272 | -34.7684 | -23.7155
14.8334 | -34.7689 | -23.7234
14.8395 | -34.7694 | -23.7313
14.8456 | -34.7698 | -23.7391
14.8517 | -34.7703 | -23.7472
14.8578 | -34.7707 | -23.755
14.8639 | -34.7712 | -23.7632
14.87 | -34.7717 | -23.771
14.8761 | -34.7721 | -23.7788
14.8822 | -34.7725 | -23.7868
14.8883 | -34.773 | -23.7945
14.8944 | -34.7735 | -23.8025
14.9005 | -34.7739 | -23.8101
14.9066 | -34.7743 | -23.818
14.9127 | -34.7748 | -23.8258
14.9188 | -34.7753 | -23.834
14.9249 | -34.7757 | -23.8419
14.931 | -34.7761 | -23.8497
14.9371 | -34.7766 | -23.8576
14.9432 | -34.7771 | -23.8655
14.9493 | -34.7775 | -23.8733
14.9554 | -34.778 | -23.8811
14.9615 | -34.7784 | -23.889
14.9676 | -34.7789 | -23.8969
14.9738 | -34.7793 | -23.9047
14.9799 | -34.7798 | -23.9129
14.986 | -34.7803 | -23.9203
14.9921 | -34.7808 | -23.928
14.9982 | -34.7812 | -23.9363
15.0043 | -34.7816 | -23.9442
15.0104 | -34.7821 | -23.952
15.0165 | -34.7826 | -23.96
15.0226 | -34.783 | -23.9676
15.0287 | -34.7835 | -23.9759
15.0348 | -34.7839 | -23.9838
15.0409 | -34.7844 | -23.9919
15.047 | -34.7849 | -23.9998
15.0531 | -34.7855 | -24.0078
15.0592 | -34.7859 | -24.0159
15.0653 | -34.7864 | -24.0236
15.0714 | -34.7868 | -24.0313
15.0775 | -34.7873 | -24.0389
15.0836 | -34.7878 | -24.0468
15.0897 | -34.7883 | -24.0545
15.0958 | -34.7887 | -24.0622
15.1019 | -34.7892 | -24.07
15.108 | -34.7896 | -24.0779
15.1141 | -34.79 | -24.0858
15.1203 | -34.7905 | -24.0935
15.1264 | -34.7911 | -24.1013
15.1325 | -34.7916 | -24.1089
15.1386 | -34.792 | -24.1168
15.1447 | -34.7925 | -24.1247
15.1508 | -34.793 | -24.1326
15.1569 | -34.7935 | -24.1407
15.163 | -34.794 | -24.148
15.1691 | -34.7945 | -24.1558
15.1752 | -34.795 | -24.1637
15.1813 | -34.7954 | -24.1717
15.1874 | -34.7959 | -24.1796
15.1935 | -34.7965 | -24.187
15.1996 | -34.797 | -24.1952
15.2057 | -34.7974 | -24.2028
15.2118 | -34.7979 | -24.2106
15.2179 | -34.7983 | -24.2181
15.224 | -34.7988 | -24.2261
15.2301 | -34.7993 | -24.2336
15.2362 | -34.7998 | -24.2414
15.2423 | -34.8003 | -24.2495
15.2484 | -34.8007 | -24.2571
15.2545 | -34.8012 | -24.2648
15.2607 | -34.8017 | -24.2728
15.2668 | -34.8022 | -24.2804
15.2729 | -34.8027 | -24.2881
15.279 | -34.8031 | -24.2959
15.2851 | -34.8036 | -24.3036
15.2912 | -34.8041 | -24.3113
15.2973 | -34.8045 | -24.319
15.3034 | -34.805 | -24.3264
15.3095 | -34.8055 | -24.3339
15.3156 | -34.8059 | -24.3418
15.3217 | -34.8064 | -24.3492
15.3278 | -34.8069 | -24.357
15.3339 | -34.8074 | -24.3644
15.34 | -34.8079 | -24.372
15.3461 | -34.8083 | -24.3798
15.3522 | -34.8088 | -24.3875
15.3583 | -34.8093 | -24.3952
15.3644 | -34.8097 | -24.403
15.3705 | -34.8102 | -24.4105
15.3766 | -34.8107 | -24.4185
15.3827 | -34.8112 | -24.426
15.3888 | -34.8116 | -24.4334
15.3949 | -34.8121 | -24.4409
15.401 | -34.8126 | -24.4486
15.4072 | -34.8131 | -24.456
15.4133 | -34.8136 | -24.4638
15.4194 | -34.8141 | -24.4715
15.4255 | -34.8146 | -24.4791
15.4316 | -34.8151 | -24.4869
15.4377 | -34.8156 | -24.4944
15.4438 | -34.8161 | -24.5017
15.4499 | -34.8165 | -24.5093
15.456 | -34.817 | -24.5171
15.4621 | -34.8175 | -24.5247
15.4682 | -34.818 | -24.5324
15.4743 | -34.8185 | -24.5399
15.4804 | -34.819 | -24.5477
15.4865 | -34.8195 | -24.5553
15.4926 | -34.82 | -24.563
15.4987 | -34.8205 | -24.5707
15.5048 | -34.8209 | -24.5783
15.5109 | -34.8215 | -24.5855
15.517 | -34.822 | -24.5933
15.5231 | -34.8224 | -24.6009
15.5292 | -34.8229 | -24.6085
15.5353 | -34.8234 | -24.6161
15.5414 | -34.8239 | -24.6241
15.5476 | -34.8244 | -24.6319
15.5537 | -34.8249 | -24.6396
15.5598 | -34.8254 | -24.6469
15.5659 | -34.8259 | -24.6548
15.572 | -34.8264 | -24.6621
15.5781 | -34.8268 | -24.6698
15.5842 | -34.8274 | -24.6772
15.5903 | -34.8278 | -24.6846
15.5964 | -34.8283 | -24.692
15.6025 | -34.8288 | -24.6995
15.6086 | -34.8293 | -24.7072
15.6147 | -34.8297 | -24.7148
15.6208 | -34.8303 | -24.7222
15.6269 | -34.8308 | -24.7297
15.633 | -34.8313 | -24.7373
15.6391 | -34.8318 | -24.745
15.6452 | -34.8323 | -24.7524
15.6513 | -34.8328 | -24.7603
15.6574 | -34.8333 | -24.7676
15.6635 | -34.8337 | -24.7752
15.6696 | -34.8342 | -24.7827
15.6757 | -34.8347 | -24.7905
15.6818 | -34.8352 | -24.7984
15.688 | -34.8358 | -24.8059
15.6941 | -34.8363 | -24.8136
15.7002 | -34.8368 | -24.8212
15.7063 | -34.8372 | -24.829
15.7124 | -34.8377 | -24.8365
15.7185 | -34.8382 | -24.844
15.7246 | -34.8387 | -24.8514
15.7307 | -34.8392 | -24.8591
15.7368 | -34.8397 | -24.867
15.7429 | -34.8402 | -24.8745
15.749 | -34.8407 | -24.882
15.7551 | -34.8411 | -24.8895
15.7612 | -34.8416 | -24.8966
15.7673 | -34.8421 | -24.9042
15.7734 | -34.8425 | -24.9115
15.7795 | -34.843 | -24.9195
15.7856 | -34.8435 | -24.927
15.7917 | -34.844 | -24.9344
15.7978 | -34.8446 | -24.9418
15.8039 | -34.845 | -24.9493
15.81 | -34.8455 | -24.9567
15.8161 | -34.846 | -24.9645
15.8222 | -34.8465 | -24.9721
15.8283 | -34.847 | -24.9793
15.8345 | -34.8475 | -24.987
15.8406 | -34.848 | -24.9946
15.8467 | -34.8485 | -25.002
15.8528 | -34.849 | -25.0094
15.8589 | -34.8495 | -25.0167
15.865 | -34.85 | -25.0241
15.8711 | -34.8505 | -25.0318
15.8772 | -34.851 | -25.0395
15.8833 | -34.8515 | -25.0473
15.8894 | -34.852 | -25.0546
15.8955 | -34.8525 | -25.0624
15.9016 | -34.853 | -25.07
15.9077 | -34.8536 | -25.0775
15.9138 | -34.8541 | -25.0851
15.9199 | -34.8546 | -25.0927
15.926 | -34.8551 | -25.1006
15.9321 | -34.8556 | -25.1081
15.9382 | -34.8561 | -25.1157
15.9443 | -34.8567 | -25.1235
15.9504 | -34.8572 | -25.1308
15.9565 | -34.8577 | -25.1383
15.9626 | -34.8582 | -25.1459
15.9687 | -34.8588 | -25.1533
15.9749 | -34.8593 | -25.161
15.981 | -34.8598 | -25.1685
15.9871 | -34.8603 | -25.176
15.9932 | -34.8608 | -25.1837
15.9993 | -34.8613 | -25.1911
16.0054 | -34.8618 | -25.1985
16.0115 | -34.8624 | -25.2057
16.0176 | -34.8629 | -25.2135
16.0237 | -34.8634 | -25.2213
16.0298 | -34.8639 | -25.2285
16.0359 | -34.8644 | -25.236
16.042 | -34.8649 | -25.2434
16.0481 | -34.8654 | -25.2509
16.0542 | -34.8659 | -25.2587
16.0603 | -34.8665 | -25.2661
16.0664 | -34.867 | -25.2736
16.0725 | -34.8675 | -25.2811
16.0786 | -34.8681 | -25.2888
16.0847 | -34.8685 | -25.2962
16.0908 | -34.8691 | -25.3038
16.0969 | -34.8696 | -25.3109
16.103 | -34.87 | -25.3186
16.1091 | -34.8705 | -25.3261
16.1152 | -34.871 | -25.334
16.1214 | -34.8715 | -25.3415
16.1275 | -34.872 | -25.3488
16.1336 | -34.8725 | -25.3563
16.1397 | -34.873 | -25.3638
16.1458 | -34.8735 | -25.3712
16.1519 | -34.874 | -25.3788
16.158 | -34.8745 | -25.3864
16.1641 | -34.8751 | -25.3937
16.1702 | -34.8756 | -25.4012
16.1763 | -34.8761 | -25.4088
16.1824 | -34.8766 | -25.416
16.1885 | -34.8771 | -25.4234
16.1946 | -34.8776 | -25.4311
16.2007 | -34.8781 | -25.4392
16.2068 | -34.8786 | -25.4464
16.2129 | -34.8791 | -25.4539
16.219 | -34.8795 | -25.4611
16.2251 | -34.88 | -25.4686
16.2312 | -34.8805 | -25.4763
16.2373 | -34.881 | -25.4841
16.2434 | -34.8814 | -25.4916
16.2495 | -34.882 | -25.4987
16.2556 | -34.8825 | -25.5063
16.2618 | -34.883 | -25.5135
16.2679 | -34.8835 | -25.5209
16.274 | -34.884 | -25.5285
16.2801 | -34.8845 | -25.5358
16.2862 | -34.8851 | -25.543
16.2923 | -34.8855 | -25.5506
16.2984 | -34.886 | -25.5582
16.3045 | -34.8866 | -25.5657
16.3106 | -34.8871 | -25.5729
16.3167 | -34.8876 | -25.5804
16.3228 | -34.8882 | -25.588
16.3289 | -34.8887 | -25.5954
16.335 | -34.8892 | -25.6032
16.3411 | -34.8896 | -25.6107
16.3472 | -34.8902 | -25.6181
16.3533 | -34.8907 | -25.6257
16.3594 | -34.8911 | -25.6333
16.3655 | -34.8916 | -25.6408
16.3716 | -34.8921 | -25.6482
16.3777 | -34.8926 | -25.6556
16.3838 | -34.8931 | -25.663
16.3899 | -34.8937 | -25.6705
16.396 | -34.8941 | -25.6782
16.4021 | -34.8947 | -25.6855
16.4083 | -34.8952 | -25.693
16.4144 | -34.8957 | -25.7004
16.4205 | -34.8962 | -25.708
16.4266 | -34.8967 | -25.7152
16.4327 | -34.8973 | -25.7225
16.4388 | -34.8977 | -25.73
16.4449 | -34.8982 | -25.7374
16.451 | -34.8987 | -25.7452
16.4571 | -34.8992 | -25.753
16.4632 | -34.8997 | -25.7603
16.4693 | -34.9002 | -25.7678
16.4754 | -34.9007 | -25.7753
16.4815 | -34.9012 | -25.7829
16.4876 | -34.9017 | -25.7905
16.4937 | -34.9022 | -25.7979
16.4998 | -34.9028 | -25.8055
16.5059 | -34.9033 | -25.8129
16.512 | -34.9038 | -25.8206
16.5181 | -34.9043 | -25.8281
16.5242 | -34.9048 | -25.8356
16.5303 | -34.9053 | -25.8428
16.5364 | -34.9058 | -25.8504
16.5425 | -34.9063 | -25.8581
16.5487 | -34.9068 | -25.8655
16.5548 | -34.9073 | -25.873
16.5609 | -34.9078 | -25.8807
16.567 | -34.9082 | -25.8885
16.5731 | -34.9088 | -25.8959
16.5792 | -34.9093 | -25.9034
16.5853 | -34.9098 | -25.9106
16.5914 | -34.9103 | -25.9184
16.5975 | -34.9108 | -25.9262
16.6036 | -34.9113 | -25.9337
16.6097 | -34.9118 | -25.9412
16.6158 | -34.9123 | -25.9488
16.6219 | -34.9128 | -25.9561
16.628 | -34.9133 | -25.9638
16.6341 | -34.9138 | -25.971
16.6402 | -34.9144 | -25.9784
16.6463 | -34.9148 | -25.9859
16.6524 | -34.9153 | -25.9934
16.6585 | -34.9158 | -26.0009
16.6646 | -34.9163 | -26.0085
16.6707 | -34.9168 | -26.016
16.6768 | -34.9174 | -26.0234
16.6829 | -34.9179 | -26.031
16.689 | -34.9184 | -26.0385
16.6952 | -34.9189 | -26.046
16.7013 | -34.9194 | -26.0536
16.7074 | -34.9199 | -26.0609
16.7135 | -34.9203 | -26.0683
16.7196 | -34.9208 | -26.0757
16.7257 | -34.9213 | -26.0831
16.7318 | -34.9218 | -26.0905
16.7379 | -34.9223 | -26.0981
16.744 | -34.9228 | -26.1053
16.7501 | -34.9234 | -26.1129
16.7562 | -34.9239 | -26.1202
16.7623 | -34.9244 | -26.1279
16.7684 | -34.9249 | -26.1355
16.7745 | -34.9254 | -26.1426
16.7806 | -34.9259 | -26.1502
16.7867 | -34.9264 | -26.1576
16.7928 | -34.9269 | -26.1653
16.7989 | -34.9274 | -26.1729
16.805 | -34.9278 | -26.1805
16.8111 | -34.9284 | -26.1879
16.8172 | -34.9289 | -26.1957
16.8233 | -34.9294 | -26.2034
16.8294 | -34.9299 | -26.2109
16.8356 | -34.9304 | -26.2187
16.8417 | -34.931 | -26.2261
16.8478 | -34.9314 | -26.2337
16.8539 | -34.932 | -26.2412
16.86 | -34.9325 | -26.2485
16.8661 | -34.9329 | -26.2559
16.8722 | -34.9335 | -26.2634
16.8783 | -34.934 | -26.2707
16.8844 | -34.9345 | -26.2783
16.8905 | -34.935 | -26.286
16.8966 | -34.9355 | -26.2936
16.9027 | -34.936 | -26.3006
16.9088 | -34.9365 | -26.3086
16.9149 | -34.937 | -26.3165
16.921 | -34.9375 | -26.3241
16.9271 | -34.938 | -26.3318
16.9332 | -34.9386 | -26.3395
16.9393 | -34.9391 | -26.3475
16.9454 | -34.9395 | -26.3546
16.9515 | -34.94 | -26.3623
16.9576 | -34.9404 | -26.3697
16.9637 | -34.9409 | -26.3774
16.9698 | -34.9414 | -26.3848
16.9759 | -34.9419 | -26.3922
16.9821 | -34.9424 | -26.3999
16.9882 | -34.9429 | -26.4076
16.9943 | -34.9433 | -26.4154
17.0004 | -34.9438 | -26.4229
17.0065 | -34.9443 | -26.4304
17.0126 | -34.9448 | -26.4379
17.0187 | -34.9453 | -26.4453
17.0248 | -34.9458 | -26.4524
17.0309 | -34.9463 | -26.4599
17.037 | -34.9468 | -26.4676
17.0431 | -34.9473 | -26.4748
17.0492 | -34.9478 | -26.4822
17.0553 | -34.9483 | -26.4897
17.0614 | -34.9487 | -26.4971
17.0675 | -34.9492 | -26.5049
17.0736 | -34.9497 | -26.5126
17.0797 | -34.9502 | -26.5197
17.0858 | -34.9507 | -26.5272
17.0919 | -34.9511 | -26.535
17.098 | -34.9516 | -26.5424
17.1041 | -34.952 | -26.5498
17.1102 | -34.9525 | -26.5573
17.1163 | -34.9531 | -26.5649
17.1225 | -34.9535 | -26.5723
17.1286 | -34.954 | -26.5796
17.1347 | -34.9545 | -26.587
17.1408 | -34.955 | -26.5945
17.1469 | -34.9555 | -26.6018
17.153 | -34.956 | -26.6094
17.1591 | -34.9565 | -26.6168
17.1652 | -34.957 | -26.6243
17.1713 | -34.9575 | -26.6318
17.1774 | -34.958 | -26.6393
17.1835 | -34.9585 | -26.6471
17.1896 | -34.959 | -26.6546
17.1957 | -34.9594 | -26.662
17.2018 | -34.9599 | -26.6696
17.2079 | -34.9604 | -26.6769
17.214 | -34.9609 | -26.6844
17.2201 | -34.9614 | -26.6919
17.2262 | -34.9618 | -26.6994
17.2323 | -34.9623 | -26.7068
17.2384 | -34.9628 | -26.7142
17.2445 | -34.9633 | -26.7217
17.2506 | -34.9637 | -26.7293
17.2567 | -34.9642 | -26.7368
17.2628 | -34.9647 | -26.7447
17.269 | -34.9651 | -26.7522
17.2751 | -34.9656 | -26.7597
17.2812 | -34.9661 | -26.7669
17.2873 | -34.9666 | -26.7744
17.2934 | -34.9671 | -26.7818
17.2995 | -34.9676 | -26.7892
17.3056 | -34.968 | -26.7967
17.3117 | -34.9685 | -26.8041
17.3178 | -34.969 | -26.8116
17.3239 | -34.9695 | -26.8189
17.33 | -34.97 | -26.8264
17.3361 | -34.9704 | -26.8341
17.3422 | -34.9709 | -26.8416
17.3483 | -34.9714 | -26.8491
17.3544 | -34.9719 | -26.8565
17.3605 | -34.9724 | -26.8641
17.3666 | -34.9729 | -26.8715
17.3727 | -34.9734 | -26.879
17.3788 | -34.9738 | -26.8862
17.3849 | -34.9743 | -26.8938
17.391 | -34.9748 | -26.9011
17.3971 | -34.9753 | -26.9087
17.4032 | -34.9758 | -26.9165
17.4094 | -34.9763 | -26.924
17.4155 | -34.9767 | -26.9313
17.4216 | -34.9772 | -26.9389
17.4277 | -34.9777 | -26.9465
17.4338 | -34.9782 | -26.954
17.4399 | -34.9787 | -26.9615
17.446 | -34.9792 | -26.9692
17.4521 | -34.9796 | -26.9764
17.4582 | -34.9801 | -26.9838
17.4643 | -34.9806 | -26.9915
17.4704 | -34.981 | -26.9989
17.4765 | -34.9815 | -27.0064
17.4826 | -34.982 | -27.0136
17.4887 | -34.9825 | -27.0214
17.4948 | -34.983 | -27.0287
17.5009 | -34.9834 | -27.036
17.507 | -34.9839 | -27.0435
17.5131 | -34.9844 | -27.0508
17.5192 | -34.9849 | -27.0583
17.5253 | -34.9853 | -27.0659
17.5314 | -34.9858 | -27.0733
17.5375 | -34.9863 | -27.0809
17.5436 | -34.9867 | -27.0883
17.5497 | -34.9872 | -27.0959
17.5559 | -34.9876 | -27.1032
17.562 | -34.9881 | -27.1104
17.5681 | -34.9886 | -27.1178
17.5742 | -34.9891 | -27.1253
17.5803 | -34.9895 | -27.1326
17.5864 | -34.99 | -27.1398
17.5925 | -34.9905 | -27.1475
17.5986 | -34.991 | -27.1547
17.6047 | -34.9915 | -27.1623
17.6108 | -34.992 | -27.1696
17.6169 | -34.9924 | -27.177
17.623 | -34.9929 | -27.1842
17.6291 | -34.9934 | -27.1915
17.6352 | -34.9939 | -27.1988
17.6413 | -34.9943 | -27.2065
17.6474 | -34.9948 | -27.2139
17.6535 | -34.9953 | -27.2216
17.6596 | -34.9957 | -27.2291
17.6657 | -34.9962 | -27.2364
17.6718 | -34.9967 | -27.244
17.6779 | -34.9971 | -27.2514
17.684 | -34.9976 | -27.2589
17.6901 | -34.9981 | -27.2662
17.6963 | -34.9986 | -27.2736
17.7024 | -34.9991 | -27.281
17.7085 | -34.9995 | -27.2883
17.7146 | -35 | -27.2957
17.7207 | -35.0005 | -27.3029
17.7268 | -35.001 | -27.3102
17.7329 | -35.0014 | -27.3179
17.739 | -35.0019 | -27.3253
17.7451 | -35.0024 | -27.3328
17.7512 | -35.0028 | -27.3402
17.7573 | -35.0033 | -27.3476
17.7634 | -35.0038 | -27.3547
17.7695 | -35.0042 | -27.3621
17.7756 | -35.0047 | -27.3696
17.7817 | -35.0052 | -27.3767
17.7878 | -35.0057 | -27.3839
17.7939 | -35.0062 | -27.3911
17.8 | -35.0067 | -27.3986
17.8061 | -35.0071 | -27.4062
17.8122 | -35.0076 | -27.4136
17.8183 | -35.008 | -27.4212
17.8244 | -35.0085 | -27.4285
17.8305 | -35.009 | -27.4359
17.8366 | -35.0095 | -27.4431
17.8428 | -35.0099 | -27.4507
17.8489 | -35.0104 | -27.4579
17.855 | -35.0109 | -27.465
17.8611 | -35.0113 | -27.4721
17.8672 | -35.0118 | -27.4796
17.8733 | -35.0123 | -27.4872
17.8794 | -35.0127 | -27.4948
17.8855 | -35.0132 | -27.502
17.8916 | -35.0137 | -27.5093
17.8977 | -35.0142 | -27.5166
17.9038 | -35.0146 | -27.5238
17.9099 | -35.0151 | -27.5312
17.916 | -35.0156 | -27.5384
17.9221 | -35.0161 | -27.546
17.9282 | -35.0165 | -27.5534
17.9343 | -35.017 | -27.5608
17.9404 | -35.0175 | -27.5682
17.9465 | -35.0179 | -27.5757
17.9526 | -35.0184 | -27.5832
17.9587 | -35.0188 | -27.5906
17.9648 | -35.0193 | -27.5978
17.9709 | -35.0198 | -27.6049
17.977 | -35.0203 | -27.612
17.9832 | -35.0208 | -27.6195
17.9893 | -35.0212 | -27.6269
17.9954 | -35.0217 | -27.6342
18.0015 | -35.0222 | -27.6415
18.0076 | -35.0226 | -27.6485
18.0137 | -35.0231 | -27.6562
18.0198 | -35.0235 | -27.6637
18.0259 | -35.024 | -27.6712
18.032 | -35.0244 | -27.6782
18.0381 | -35.0249 | -27.6854
18.0442 | -35.0254 | -27.6928
18.0503 | -35.0258 | -27.7003
18.0564 | -35.0262 | -27.7074
18.0625 | -35.0267 | -27.7148
18.0686 | -35.0272 | -27.722
18.0747 | -35.0277 | -27.7291
18.0808 | -35.0281 | -27.7361
18.0869 | -35.0286 | -27.7438
18.093 | -35.0291 | -27.7513
18.0991 | -35.0296 | -27.7589
18.1052 | -35.0301 | -27.7663
18.1113 | -35.0305 | -27.7732
18.1174 | -35.031 | -27.7805
18.1236 | -35.0315 | -27.788
18.1297 | -35.032 | -27.7952
18.1358 | -35.0325 | -27.8026
18.1419 | -35.033 | -27.8101
18.148 | -35.0335 | -27.8178
18.1541 | -35.034 | -27.8247
18.1602 | -35.0345 | -27.832
18.1663 | -35.035 | -27.8393
18.1724 | -35.0354 | -27.8471
18.1785 | -35.0359 | -27.8544
18.1846 | -35.0363 | -27.8615
18.1907 | -35.0368 | -27.869
18.1968 | -35.0372 | -27.8763
18.2029 | -35.0377 | -27.8835
18.209 | -35.0382 | -27.8911
18.2151 | -35.0387 | -27.8986
18.2212 | -35.0392 | -27.9056
18.2273 | -35.0398 | -27.913
18.2334 | -35.0403 | -27.92
18.2395 | -35.0407 | -27.9274
18.2456 | -35.0412 | -27.9346
18.2517 | -35.0416 | -27.9418
18.2578 | -35.042 | -27.9488
18.2639 | -35.0425 | -27.9564
18.2701 | -35.043 | -27.9637
18.2762 | -35.0435 | -27.9712
18.2823 | -35.044 | -27.9786
18.2884 | -35.0444 | -27.9861
18.2945 | -35.0449 | -27.993
18.3006 | -35.0453 | -28.0003
18.3067 | -35.0458 | -28.0076
18.3128 | -35.0463 | -28.015
18.3189 | -35.0468 | -28.0223
18.325 | -35.0472 | -28.0291
18.3311 | -35.0477 | -28.0366
18.3372 | -35.0482 | -28.0437
18.3433 | -35.0486 | -28.0509
18.3494 | -35.049 | -28.0583
18.3555 | -35.0495 | -28.0653
18.3616 | -35.05 | -28.0724
18.3677 | -35.0504 | -28.0793
18.3738 | -35.0509 | -28.0864
18.3799 | -35.0513 | -28.0935
18.386 | -35.0518 | -28.1007
18.3921 | -35.0523 | -28.1079
18.3982 | -35.0527 | -28.1151
18.4043 | -35.0532 | -28.1223
18.4105 | -35.0536 | -28.1295
18.4166 | -35.0541 | -28.1366
18.4227 | -35.0546 | -28.1438
18.4288 | -35.0551 | -28.1513
18.4349 | -35.0556 | -28.1582
18.441 | -35.0561 | -28.1654
18.4471 | -35.0566 | -28.1726
18.4532 | -35.057 | -28.1798
18.4593 | -35.0575 | -28.187
18.4654 | -35.0579 | -28.1937
18.4715 | -35.0584 | -28.201
18.4776 | -35.0589 | -28.2083
18.4837 | -35.0594 | -28.2156
18.4898 | -35.0599 | -28.2231
18.4959 | -35.0604 | -28.2304
18.502 | -35.0609 | -28.2377
18.5081 | -35.0614 | -28.245
18.5142 | -35.0618 | -28.2524
18.5203 | -35.0622 | -28.2592
18.5264 | -35.0627 | -28.2661
18.5325 | -35.0632 | -28.2736
18.5386 | -35.0637 | -28.2808
18.5447 | -35.0642 | -28.288
18.5508 | -35.0648 | -28.2952
18.557 | -35.0652 | -28.3025
18.5631 | -35.0657 | -28.3098
18.5692 | -35.0661 | -28.3171
18.5753 | -35.0666 | -28.3241
18.5814 | -35.0671 | -28.3315
18.5875 | -35.0676 | -28.3391
18.5936 | -35.068 | -28.3465
18.5997 | -35.0685 | -28.3537
18.6058 | -35.069 | -28.3607
18.6119 | -35.0695 | -28.368
18.618 | -35.07 | -28.3752
18.6241 | -35.0705 | -28.3828
18.6302 | -35.071 | -28.3903
18.6363 | -35.0715 | -28.3973
18.6424 | -35.072 | -28.4043
18.6485 | -35.0725 | -28.4115
18.6546 | -35.0729 | -28.4187
18.6607 | -35.0734 | -28.4258
18.6668 | -35.0739 | -28.4328
18.6729 | -35.0743 | -28.4402
18.679 | -35.0748 | -28.4471
18.6851 | -35.0753 | -28.4543
18.6912 | -35.0758 | -28.4615
18.6974 | -35.0763 | -28.4687
18.7035 | -35.0768 | -28.4759
18.7096 | -35.0773 | -28.4831
18.7157 | -35.0778 | -28.4898
18.7218 | -35.0783 | -28.4969
18.7279 | -35.0788 | -28.5038
18.734 | -35.0793 | -28.5109
18.7401 | -35.0798 | -28.5177
18.7462 | -35.0803 | -28.5242
18.7523 | -35.0808 | -28.5314
18.7584 | -35.0813 | -28.5387
18.7645 | -35.0818 | -28.546
18.7706 | -35.0822 | -28.5531
18.7767 | -35.0827 | -28.5601
18.7828 | -35.0831 | -28.5667
18.7889 | -35.0836 | -28.5737
18.795 | -35.0841 | -28.5808
18.8011 | -35.0846 | -28.5879
18.8072 | -35.0851 | -28.5947
18.8133 | -35.0855 | -28.602
18.8194 | -35.086 | -28.6093
18.8255 | -35.0865 | -28.6162
18.8316 | -35.087 | -28.6234
18.8377 | -35.0874 | -28.6302
18.8439 | -35.0879 | -28.6372
18.85 | -35.0884 | -28.6442
18.8561 | -35.0889 | -28.6508
18.8622 | -35.0894 | -28.658
18.8683 | -35.0899 | -28.6649
18.8744 | -35.0904 | -28.6718
18.8805 | -35.0909 | -28.6793
18.8866 | -35.0914 | -28.6866
18.8927 | -35.0919 | -28.6934
18.8988 | -35.0924 | -28.7002
18.9049 | -35.0929 | -28.7073
18.911 | -35.0934 | -28.7143
18.9171 | -35.0939 | -28.7214
18.9232 | -35.0944 | -28.728
18.9293 | -35.0948 | -28.7351
18.9354 | -35.0953 | -28.7421
18.9415 | -35.0957 | -28.7491
18.9476 | -35.0962 | -28.7564
18.9537 | -35.0967 | -28.7637
18.9598 | -35.0972 | -28.7704
18.9659 | -35.0977 | -28.7774
18.972 | -35.0982 | -28.7844
18.9781 | -35.0987 | -28.791
18.9843 | -35.0992 | -28.798
18.9904 | -35.0997 | -28.805
18.9965 | -35.1001 | -28.8121
19.0026 | -35.1006 | -28.8191
19.0087 | -35.1011 | -28.826
19.0148 | -35.1016 | -28.8325
19.0209 | -35.102 | -28.8397
19.027 | -35.1025 | -28.8468
19.0331 | -35.103 | -28.8536
19.0392 | -35.1035 | -28.8607
19.0453 | -35.104 | -28.8677
19.0514 | -35.1044 | -28.8748
19.0575 | -35.1049 | -28.8821
19.0636 | -35.1055 | -28.8887
19.0697 | -35.106 | -28.896
19.0758 | -35.1065 | -28.9026
19.0819 | -35.107 | -28.9096
19.088 | -35.1074 | -28.9162
19.0941 | -35.1079 | -28.923
19.1002 | -35.1084 | -28.9298
19.1063 | -35.1089 | -28.9368
19.1124 | -35.1094 | -28.9439
19.1185 | -35.1099 | -28.9515
19.1246 | -35.1103 | -28.9584
19.1308 | -35.1108 | -28.9651
19.1369 | -35.1114 | -28.9716
19.143 | -35.1118 | -28.9786
19.1491 | -35.1123 | -28.9855
19.1552 | -35.1128 | -28.993
19.1613 | -35.1132 | -29
19.1674 | -35.1137 | -29.007
19.1735 | -35.1142 | -29.0138
19.1796 | -35.1147 | -29.0206
19.1857 | -35.1152 | -29.0278
19.1918 | -35.1157 | -29.0347
19.1979 | -35.1162 | -29.0419
19.204 | -35.1167 | -29.0486
19.2101 | -35.1172 | -29.0556
19.2162 | -35.1176 | -29.0631
19.2223 | -35.1181 | -29.0695
19.2284 | -35.1187 | -29.0767
19.2345 | -35.1192 | -29.0836
19.2406 | -35.1197 | -29.0907
19.2467 | -35.1201 | -29.0975
19.2528 | -35.1206 | -29.1042
19.2589 | -35.121 | -29.1113
19.265 | -35.1215 | -29.1184
19.2712 | -35.122 | -29.1255
19.2773 | -35.1225 | -29.1322
19.2834 | -35.123 | -29.1391
19.2895 | -35.1235 | -29.1461
19.2956 | -35.1239 | -29.1533
19.3017 | -35.1244 | -29.1602
19.3078 | -35.1249 | -29.1672
19.3139 | -35.1254 | -29.1739
19.32 | -35.1259 | -29.1807
19.3261 | -35.1263 | -29.1876
19.3322 | -35.1268 | -29.1945
19.3383 | -35.1274 | -29.2011
19.3444 | -35.1278 | -29.2076
19.3505 | -35.1283 | -29.2143
19.3566 | -35.1288 | -29.2212
19.3627 | -35.1293 | -29.2281
19.3688 | -35.1298 | -29.2353
19.3749 | -35.1303 | -29.2424
19.381 | -35.1307 | -29.2486
19.3871 | -35.1312 | -29.2551
19.3932 | -35.1317 | -29.262
19.3993 | -35.1321 | -29.2687
19.4054 | -35.1326 | -29.276
19.4115 | -35.1331 | -29.2828
19.4177 | -35.1336 | -29.2898
19.4238 | -35.1341 | -29.2967
19.4299 | -35.1346 | -29.3035
19.436 | -35.1351 | -29.3109
19.4421 | -35.1356 | -29.3172
19.4482 | -35.1362 | -29.3241
19.4543 | -35.1367 | -29.3306
19.4604 | -35.1372 | -29.3373
19.4665 | -35.1378 | -29.3444
19.4726 | -35.1383 | -29.351
19.4787 | -35.1387 | -29.3579
19.4848 | -35.1392 | -29.3649
19.4909 | -35.1397 | -29.3719
19.497 | -35.1402 | -29.3788
19.5031 | -35.1407 | -29.3857
19.5092 | -35.1411 | -29.3925
19.5153 | -35.1416 | -29.3996
19.5214 | -35.1421 | -29.4061
19.5275 | -35.1426 | -29.4129
19.5336 | -35.1431 | -29.4199
19.5397 | -35.1437 | -29.4268
19.5458 | -35.1442 | -29.4335
19.5519 | -35.1447 | -29.4406
19.5581 | -35.1452 | -29.4465
19.5642 | -35.1456 | -29.4536
19.5703 | -35.1461 | -29.4603
19.5764 | -35.1466 | -29.4673
19.5825 | -35.1471 | -29.4742
19.5886 | -35.1476 | -29.4807
19.5947 | -35.1481 | -29.487
19.6008 | -35.1485 | -29.4943
19.6069 | -35.149 | -29.5011
19.613 | -35.1496 | -29.5082
19.6191 | -35.1501 | -29.5151
19.6252 | -35.1505 | -29.5222
19.6313 | -35.1511 | -29.5286
19.6374 | -35.1516 | -29.5355
19.6435 | -35.1521 | -29.5425
19.6496 | -35.1525 | -29.5491
19.6557 | -35.153 | -29.5557
19.6618 | -35.1535 | -29.5623
19.6679 | -35.154 | -29.5691
19.674 | -35.1545 | -29.5759
19.6801 | -35.155 | -29.5826
19.6862 | -35.1555 | -29.5894
19.6923 | -35.1559 | -29.5963
19.6984 | -35.1563 | -29.6027
19.7046 | -35.1569 | -29.6093
19.7107 | -35.1574 | -29.6159
19.7168 | -35.1579 | -29.6227
19.7229 | -35.1583 | -29.6293
19.729 | -35.1588 | -29.6362
19.7351 | -35.1593 | -29.6426
19.7412 | -35.1598 | -29.649
19.7473 | -35.1604 | -29.6556
19.7534 | -35.1609 | -29.6627
19.7595 | -35.1614 | -29.6696
19.7656 | -35.1618 | -29.6764
19.7717 | -35.1623 | -29.6832
19.7778 | -35.1628 | -29.6904
19.7839 | -35.1633 | -29.6969
19.79 | -35.1638 | -29.7038
19.7961 | -35.1642 | -29.711
19.8022 | -35.1647 | -29.7173
19.8083 | -35.1652 | -29.7243
19.8144 | -35.1657 | -29.7311
19.8205 | -35.1662 | -29.7381
19.8266 | -35.1667 | -29.7449
19.8327 | -35.1671 | -29.7519
19.8388 | -35.1677 | -29.7587
19.845 | -35.1682 | -29.7659
19.8511 | -35.1686 | -29.7728
19.8572 | -35.1691 | -29.7797
19.8633 | -35.1696 | -29.7863
19.8694 | -35.1702 | -29.7931
19.8755 | -35.1708 | -29.7998
19.8816 | -35.1713 | -29.8069
19.8877 | -35.1718 | -29.8137
19.8938 | -35.1722 | -29.8204
19.8999 | -35.1727 | -29.8269
19.906 | -35.1731 | -29.8337
19.9121 | -35.1736 | -29.8403
19.9182 | -35.1741 | -29.8467
19.9243 | -35.1746 | -29.8533
19.9304 | -35.1751 | -29.8604
19.9365 | -35.1755 | -29.8673
19.9426 | -35.176 | -29.8739
19.9487 | -35.1766 | -29.8807
19.9548 | -35.177 | -29.8879
19.9609 | -35.1776 | -29.8946
19.967 | -35.1782 | -29.9013
19.9731 | -35.1786 | -29.9084
19.9792 | -35.1791 | -29.9146
19.9853 | -35.1796 | -29.9213
19.9915 | -35.1802 | -29.9278
19.9976 | -35.1807 | -29.9344
20.0037 | -35.1812 | -29.9409
20.0098 | -35.1817 | -29.9476
20.0159 | -35.1822 | -29.9543
20.022 | -35.1827 | -29.9612
20.0281 | -35.1832 | -29.9681
20.0342 | -35.1837 | -29.9748
20.0403 | -35.1842 | -29.9811
20.0464 | -35.1846 | -29.9883
20.0525 | -35.1852 | -29.9946
20.0586 | -35.1856 | -30.0009
20.0647 | -35.1861 | -30.0078
20.0708 | -35.1866 | -30.0147
20.0769 | -35.1871 | -30.0215
20.083 | -35.1876 | -30.0277
20.0891 | -35.1881 | -30.0347
20.0952 | -35.1887 | -30.0414
20.1013 | -35.1891 | -30.0481
20.1074 | -35.1896 | -30.0549
20.1135 | -35.1901 | -30.0613
20.1196 | -35.1906 | -30.0683
20.1257 | -35.1911 | -30.0755
20.1319 | -35.1915 | -30.0825
20.138 | -35.192 | -30.0892
20.1441 | -35.1924 | -30.0958
20.1502 | -35.1929 | -30.1027
20.1563 | -35.1933 | -30.1092
20.1624 | -35.1937 | -30.116
20.1685 | -35.1943 | -30.1226
20.1746 | -35.1948 | -30.129
20.1807 | -35.1952 | -30.136
20.1868 | -35.1957 | -30.1427
20.1929 | -35.1962 | -30.1494
20.199 | -35.1967 | -30.1559
20.2051 | -35.1972 | -30.1628
20.2112 | -35.1977 | -30.1688
20.2173 | -35.1981 | -30.1754
20.2234 | -35.1987 | -30.1821
20.2295 | -35.1992 | -30.1891
20.2356 | -35.1997 | -30.196
20.2417 | -35.2003 | -30.2027
20.2478 | -35.2007 | -30.2095
20.2539 | -35.2011 | -30.2165
20.26 | -35.2016 | -30.2232
20.2661 | -35.202 | -30.2297
20.2723 | -35.2025 | -30.2361
20.2784 | -35.203 | -30.2428
20.2845 | -35.2035 | -30.2501
20.2906 | -35.204 | -30.2576
20.2967 | -35.2045 | -30.2643
20.3028 | -35.2051 | -30.2707
20.3089 | -35.2055 | -30.2773
20.315 | -35.206 | -30.2838
20.3211 | -35.2065 | -30.2909
20.3272 | -35.2071 | -30.2975
20.3333 | -35.2075 | -30.3039
20.3394 | -35.2079 | -30.31
20.3455 | -35.2083 | -30.3169
20.3516 | -35.2088 | -30.324
20.3577 | -35.2092 | -30.3301
20.3638 | -35.2097 | -30.3367
20.3699 | -35.2101 | -30.3426
20.376 | -35.2106 | -30.3487
20.3821 | -35.2111 | -30.3544
20.3882 | -35.2116 | -30.3614
20.3943 | -35.2121 | -30.3681
20.4004 | -35.2127 | -30.3739
20.4065 | -35.2133 | -30.3807
20.4126 | -35.2137 | -30.3872
20.4188 | -35.2142 | -30.3943
20.4249 | -35.2147 | -30.401
20.431 | -35.2152 | -30.407
20.4371 | -35.2157 | -30.4127
20.4432 | -35.2162 | -30.4189
20.4493 | -35.2167 | -30.4261
20.4554 | -35.2171 | -30.4323
20.4615 | -35.2176 | -30.4391
20.4676 | -35.2181 | -30.4458
20.4737 | -35.2185 | -30.4521
20.4798 | -35.219 | -30.4593
20.4859 | -35.2196 | -30.4654
20.492 | -35.2201 | -30.4719
20.4981 | -35.2205 | -30.4786
20.5042 | -35.221 | -30.4848
20.5103 | -35.2215 | -30.4908
20.5164 | -35.222 | -30.4977
20.5225 | -35.2224 | -30.5046
20.5286 | -35.223 | -30.5113
20.5347 | -35.2234 | -30.5175
20.5408 | -35.224 | -30.5239
20.5469 | -35.2245 | -30.5301
20.553 | -35.225 | -30.5365
20.5592 | -35.2255 | -30.5423
20.5653 | -35.226 | -30.5491
20.5714 | -35.2264 | -30.5556
20.5775 | -35.2269 | -30.5616
20.5836 | -35.2274 | -30.5681
20.5897 | -35.228 | -30.5745
20.5958 | -35.2285 | -30.5813
20.6019 | -35.229 | -30.5882
20.608 | -35.2294 | -30.5944
20.6141 | -35.2299 | -30.6008
20.6202 | -35.2303 | -30.6074
20.6263 | -35.2307 | -30.6136
20.6324 | -35.2312 | -30.6198
20.6385 | -35.2316 | -30.6261
20.6446 | -35.2321 | -30.6328
20.6507 | -35.2325 | -30.6388
20.6568 | -35.233 | -30.6453
20.6629 | -35.2335 | -30.6516
20.669 | -35.2339 | -30.6579
20.6751 | -35.2344 | -30.665
20.6812 | -35.2349 | -30.6707
20.6873 | -35.2353 | -30.6768
20.6934 | -35.2358 | -30.6832
20.6995 | -35.2363 | -30.6895
20.7057 | -35.2368 | -30.6961
20.7118 | -35.2373 | -30.7028
20.7179 | -35.2377 | -30.7091
20.724 | -35.2382 | -30.7153
20.7301 | -35.2386 | -30.7226
20.7362 | -35.239 | -30.7288
20.7423 | -35.2396 | -30.7348
20.7484 | -35.2401 | -30.7413
20.7545 | -35.2406 | -30.7474
20.7606 | -35.2411 | -30.7543
20.7667 | -35.2414 | -30.7608
20.7728 | -35.2419 | -30.7675
20.7789 | -35.2424 | -30.7739
20.785 | -35.2429 | -30.7804
20.7911 | -35.2434 | -30.7856
20.7972 | -35.2439 | -30.7918
20.8033 | -35.2444 | -30.7978
20.8094 | -35.2449 | -30.8044
20.8155 | -35.2453 | -30.811
20.8216 | -35.2458 | -30.8178
20.8277 | -35.2461 | -30.8233
20.8338 | -35.2466 | -30.8301
20.8399 | -35.2471 | -30.8366
20.8461 | -35.2476 | -30.8433
20.8522 | -35.2481 | -30.8493
20.8583 | -35.2487 | -30.8562
20.8644 | -35.2491 | -30.863
20.8705 | -35.2496 | -30.8696
20.8766 | -35.25 | -30.8764
20.8827 | -35.2506 | -30.8827
20.8888 | -35.251 | -30.8895
20.8949 | -35.2515 | -30.8965
20.901 | -35.252 | -30.9028
20.9071 | -35.2524 | -30.9091
20.9132 | -35.2528 | -30.9154
20.9193 | -35.2533 | -30.9224
20.9254 | -35.2538 | -30.9283
20.9315 | -35.2541 | -30.9341
20.9376 | -35.2546 | -30.9409
20.9437 | -35.2551 | -30.9475
20.9498 | -35.2556 | -30.9539
20.9559 | -35.2561 | -30.9598
20.962 | -35.2566 | -30.9667
20.9681 | -35.257 | -30.9732
20.9742 | -35.2575 | -30.9789
20.9803 | -35.2579 | -30.9857
20.9864 | -35.2584 | -30.992
20.9926 | -35.2589 | -30.9982
20.9987 | -35.2594 | -31.0037
21.0048 | -35.2599 | -31.0102
21.0109 | -35.2603 | -31.0165
21.017 | -35.2609 | -31.0224
21.0231 | -35.2614 | -31.0285
21.0292 | -35.2619 | -31.0353
21.0353 | -35.2625 | -31.0411
21.0414 | -35.2629 | -31.0478
21.0475 | -35.2635 | -31.0544
21.0536 | -35.264 | -31.0603
21.0597 | -35.2645 | -31.0665
21.0658 | -35.265 | -31.0723
21.0719 | -35.2655 | -31.0784
21.078 | -35.266 | -31.085
21.0841 | -35.2664 | -31.0912
21.0902 | -35.2669 | -31.0979
21.0963 | -35.2674 | -31.1042
21.1024 | -35.2678 | -31.1104
21.1085 | -35.2683 | -31.1167
21.1146 | -35.2688 | -31.1232
21.1207 | -35.2694 | -31.1295
21.1268 | -35.2698 | -31.1358
21.133 | -35.2703 | -31.142
21.1391 | -35.2709 | -31.1477
21.1452 | -35.2714 | -31.1543
21.1513 | -35.2719 | -31.1609
21.1574 | -35.2724 | -31.1678
21.1635 | -35.2727 | -31.1738
21.1696 | -35.2732 | -31.1794
21.1757 | -35.2738 | -31.1851
21.1818 | -35.2743 | -31.1916
21.1879 | -35.2748 | -31.1977
21.194 | -35.2752 | -31.2042
21.2001 | -35.2757 | -31.2104
21.2062 | -35.2762 | -31.2158
21.2123 | -35.2766 | -31.2209
21.2184 | -35.2772 | -31.228
21.2245 | -35.2775 | -31.2347
21.2306 | -35.2781 | -31.2418
21.2367 | -35.2786 | -31.2475
21.2428 | -35.2789 | -31.2531
21.2489 | -35.2794 | -31.2592
21.255 | -35.2799 | -31.2654
21.2611 | -35.2803 | -31.2722
21.2672 | -35.2807 | -31.2786
21.2733 | -35.2812 | -31.2842
21.2795 | -35.2817 | -31.2905
21.2856 | -35.2821 | -31.2971
21.2917 | -35.2825 | -31.3027
21.2978 | -35.2829 | -31.3092
21.3039 | -35.2833 | -31.3157
21.31 | -35.2838 | -31.3216
21.3161 | -35.2843 | -31.3273
21.3222 | -35.2848 | -31.3321
21.3283 | -35.2852 | -31.3378
21.3344 | -35.2856 | -31.3436
21.3405 | -35.2861 | -31.3495
21.3466 | -35.2867 | -31.3557
21.3527 | -35.2869 | -31.3618
21.3588 | -35.2874 | -31.368
21.3649 | -35.2879 | -31.3745
21.371 | -35.2883 | -31.3804
21.3771 | -35.2888 | -31.386
21.3832 | -35.2892 | -31.3919
21.3893 | -35.2896 | -31.3972
21.3954 | -35.29 | -31.4033
21.4015 | -35.2905 | -31.4095
21.4076 | -35.291 | -31.4155
21.4137 | -35.2915 | -31.4223
21.4199 | -35.2919 | -31.4284
21.426 | -35.2923 | -31.4345
21.4321 | -35.2927 | -31.4403
21.4382 | -35.2931 | -31.4464
21.4443 | -35.2935 | -31.453
21.4504 | -35.2939 | -31.4589
21.4565 | -35.2943 | -31.4645
21.4626 | -35.2949 | -31.4707
21.4687 | -35.2952 | -31.4767
21.4748 | -35.2957 | -31.4824
21.4809 | -35.2961 | -31.4888
21.487 | -35.2965 | -31.4943
21.4931 | -35.297 | -31.4994
21.4992 | -35.2974 | -31.505
21.5053 | -35.2978 | -31.5118
21.5114 | -35.2983 | -31.5171
21.5175 | -35.2986 | -31.5228
21.5236 | -35.299 | -31.5287
21.5297 | -35.2994 | -31.5346
21.5358 | -35.2999 | -31.541
21.5419 | -35.3003 | -31.5478
21.548 | -35.3008 | -31.5542
21.5541 | -35.3012 | -31.5602
21.5602 | -35.3016 | -31.5663
21.5664 | -35.3022 | -31.572
21.5725 | -35.3026 | -31.5779
21.5786 | -35.3031 | -31.5842
21.5847 | -35.3034 | -31.5905
21.5908 | -35.3039 | -31.5969
21.5969 | -35.3043 | -31.6021
21.603 | -35.3048 | -31.6081
21.6091 | -35.3053 | -31.6146
21.6152 | -35.3057 | -31.621
21.6213 | -35.3062 | -31.6274
21.6274 | -35.3067 | -31.6335
21.6335 | -35.3071 | -31.6393
21.6396 | -35.3075 | -31.6455
21.6457 | -35.308 | -31.6514
21.6518 | -35.3084 | -31.6575
21.6579 | -35.3089 | -31.6635
21.664 | -35.3093 | -31.6701
21.6701 | -35.3096 | -31.6767
21.6762 | -35.3101 | -31.6827
21.6823 | -35.3104 | -31.6882
21.6884 | -35.3109 | -31.6943
21.6945 | -35.3112 | -31.7002
21.7006 | -35.3116 | -31.7062
21.7068 | -35.3119 | -31.7122
21.7129 | -35.3124 | -31.7186
21.719 | -35.3128 | -31.7246
21.7251 | -35.3132 | -31.7308
21.7312 | -35.3137 | -31.7361
21.7373 | -35.3142 | -31.7424
21.7434 | -35.3145 | -31.7488
21.7495 | -35.3149 | -31.7548
21.7556 | -35.3153 | -31.7607
21.7617 | -35.3157 | -31.767
21.7678 | -35.3162 | -31.7735
21.7739 | -35.3165 | -31.779
21.78 | -35.3168 | -31.7853
21.7861 | -35.3173 | -31.7907
21.7922 | -35.3178 | -31.7968
21.7983 | -35.3182 | -31.8029
21.8044 | -35.3186 | -31.8084
21.8105 | -35.3191 | -31.8158
21.8166 | -35.3195 | -31.823
21.8227 | -35.32 | -31.8289
21.8288 | -35.3204 | -31.835
21.8349 | -35.321 | -31.8407
21.841 | -35.3213 | -31.8467
21.8471 | -35.3219 | -31.8532
21.8533 | -35.3223 | -31.859
21.8594 | -35.3227 | -31.8654
21.8655 | -35.3231 | -31.8712
21.8716 | -35.3235 | -31.878
21.8777 | -35.324 | -31.8846
21.8838 | -35.3244 | -31.8908
21.8899 | -35.3248 | -31.8974
21.896 | -35.3253 | -31.903
21.9021 | -35.3258 | -31.9093
21.9082 | -35.3262 | -31.9154
21.9143 | -35.3266 | -31.9213
21.9204 | -35.327 | -31.9274
21.9265 | -35.3274 | -31.9333
21.9326 | -35.3279 | -31.9395
21.9387 | -35.3283 | -31.9457
21.9448 | -35.3288 | -31.9514
21.9509 | -35.3292 | -31.9573
21.957 | -35.3297 | -31.9634
21.9631 | -35.3301 | -31.9696
21.9692 | -35.3304 | -31.9761
21.9753 | -35.3308 | -31.9825
21.9814 | -35.3311 | -31.9881
21.9875 | -35.3314 | -31.9948
21.9937 | -35.3319 | -32.0006
21.9998 | -35.3325 | -32.0066
22.0059 | -35.3328 | -32.0131
22.012 | -35.3333 | -32.0194
22.0181 | -35.3338 | -32.0253
22.0242 | -35.3341 | -32.0316
22.0303 | -35.3347 | -32.0376
22.0364 | -35.335 | -32.0435
22.0425 | -35.3355 | -32.0492
22.0486 | -35.336 | -32.0554
22.0547 | -35.3365 | -32.0613
22.0608 | -35.337 | -32.067
22.0669 | -35.3374 | -32.0737
22.073 | -35.3378 | -32.0791
22.0791 | -35.3382 | -32.0851
22.0852 | -35.3388 | -32.0913
22.0913 | -35.339 | -32.0989
22.0974 | -35.3395 | -32.1038
22.1035 | -35.34 | -32.1096
22.1096 | -35.3403 | -32.1151
22.1157 | -35.3408 | -32.1203
22.1218 | -35.3411 | -32.1259
22.1279 | -35.3416 | -32.1298
22.134 | -35.342 | -32.1362
22.1402 | -35.3424 | -32.142
22.1463 | -35.3429 | -32.1479
22.1524 | -35.3433 | -32.1546
22.1585 | -35.3437 | -32.1602
22.1646 | -35.3441 | -32.1649
22.1707 | -35.3446 | -32.1708
22.1768 | -35.3452 | -32.1765
22.1829 | -35.3454 | -32.1827
22.189 | -35.3458 | -32.1884
22.1951 | -35.346 | -32.1947
22.2012 | -35.3464 | -32.2001
22.2073 | -35.3467 | -32.205
22.2134 | -35.3472 | -32.21
22.2195 | -35.3477 | -32.2158
22.2256 | -35.3481 | -32.2217
22.2317 | -35.3483 | -32.2276
22.2378 | -35.3488 | -32.2328
22.2439 | -35.3493 | -32.2389
22.25 | -35.3497 | -32.2449
22.2561 | -35.3501 | -32.2508
22.2622 | -35.3505 | -32.2574
22.2683 | -35.3509 | -32.2633
22.2744 | -35.3514 | -32.2693
22.2806 | -35.3518 | -32.2754
22.2867 | -35.3522 | -32.2818
22.2928 | -35.3527 | -32.287
22.2989 | -35.353 | -32.2926
22.305 | -35.3534 | -32.2985
22.3111 | -35.3538 | -32.3041
22.3172 | -35.3543 | -32.3103
22.3233 | -35.3548 | -32.3169
22.3294 | -35.3551 | -32.3226
22.3355 | -35.3556 | -32.3283
22.3416 | -35.356 | -32.334
22.3477 | -35.3565 | -32.3393
22.3538 | -35.3568 | -32.3454
22.3599 | -35.3572 | -32.3513
22.366 | -35.3576 | -32.3572
22.3721 | -35.358 | -32.3628
22.3782 | -35.3584 | -32.3692
22.3843 | -35.3588 | -32.3752
22.3904 | -35.3591 | -32.3815
22.3965 | -35.3595 | -32.3869
22.4026 | -35.3599 | -32.3923
22.4087 | -35.3602 | -32.3966
22.4148 | -35.3606 | -32.4028
22.4209 | -35.361 | -32.4091
22.4271 | -35.3615 | -32.4149
22.4332 | -35.3619 | -32.4211
22.4393 | -35.3623 | -32.4259
22.4454 | -35.3629 | -32.4298
22.4515 | -35.3633 | -32.4359
22.4576 | -35.3638 | -32.4415
22.4637 | -35.3641 | -32.4468
22.4698 | -35.3645 | -32.4522
22.4759 | -35.3647 | -32.4569
22.482 | -35.3652 | -32.4624
22.4881 | -35.3656 | -32.4684
22.4942 | -35.3659 | -32.474
22.5003 | -35.3664 | -32.4794
22.5064 | -35.3667 | -32.485
22.5125 | -35.367 | -32.4901
22.5186 | -35.3674 | -32.4958
22.5247 | -35.3678 | -32.5012
22.5308 | -35.3682 | -32.507
22.5369 | -35.3686 | -32.5131
22.543 | -35.3691 | -32.5189
22.5491 | -35.3694 | -32.5241
22.5552 | -35.3698 | -32.5299
22.5613 | -35.3703 | -32.536
22.5675 | -35.3707 | -32.5422
22.5736 | -35.371 | -32.5479
22.5797 | -35.3715 | -32.5533
22.5858 | -35.3719 | -32.559
22.5919 | -35.3723 | -32.5642
22.598 | -35.3728 | -32.57
22.6041 | -35.3731 | -32.5759
22.6102 | -35.3735 | -32.5819
22.6163 | -35.3739 | -32.5871
22.6224 | -35.3743 | -32.5937
22.6285 | -35.3746 | -32.5999
22.6346 | -35.375 | -32.6059
22.6407 | -35.3754 | -32.6118
22.6468 | -35.3757 | -32.6173
22.6529 | -35.3761 | -32.6226
22.659 | -35.3764 | -32.6288
22.6651 | -35.3768 | -32.6346
22.6712 | -35.3772 | -32.6405
22.6773 | -35.3778 | -32.6462
22.6834 | -35.3783 | -32.6523
22.6895 | -35.3788 | -32.6578
22.6956 | -35.3792 | -32.6631
22.7017 | -35.3796 | -32.6695
22.7079 | -35.38 | -32.6764
22.714 | -35.3804 | -32.682
22.7201 | -35.3809 | -32.6876
22.7262 | -35.3813 | -32.6932
22.7323 | -35.3817 | -32.6994
22.7384 | -35.3822 | -32.7057
22.7445 | -35.3825 | -32.7112
22.7506 | -35.3828 | -32.7173
22.7567 | -35.3832 | -32.7226
22.7628 | -35.3836 | -32.7277
22.7689 | -35.384 | -32.733
22.775 | -35.3844 | -32.7397
22.7811 | -35.3849 | -32.7452
22.7872 | -35.3852 | -32.7502
22.7933 | -35.3856 | -32.7553
22.7994 | -35.386 | -32.7614
22.8055 | -35.3865 | -32.7677
22.8116 | -35.3869 | -32.7734
22.8177 | -35.3873 | -32.779
22.8238 | -35.3876 | -32.7842
22.8299 | -35.388 | -32.7885
22.836 | -35.3884 | -32.7937
22.8421 | -35.3888 | -32.7991
22.8482 | -35.3891 | -32.8049
22.8544 | -35.3895 | -32.8114
22.8605 | -35.3899 | -32.8173
22.8666 | -35.3902 | -32.8227
22.8727 | -35.3905 | -32.8291
22.8788 | -35.3908 | -32.8345
22.8849 | -35.3911 | -32.8396
22.891 | -35.3915 | -32.8456
22.8971 | -35.392 | -32.8514
22.9032 | -35.3924 | -32.8566
22.9093 | -35.3928 | -32.8624
22.9154 | -35.3932 | -32.8683
22.9215 | -35.3935 | -32.8742
22.9276 | -35.3938 | -32.8804
22.9337 | -35.3943 | -32.8857
22.9398 | -35.3947 | -32.8916
22.9459 | -35.3951 | -32.8973
22.952 | -35.3955 | -32.9032
22.9581 | -35.3959 | -32.9091
22.9642 | -35.3962 | -32.9151
22.9703 | -35.3965 | -32.9204
22.9764 | -35.3969 | -32.9264
22.9825 | -35.3973 | -32.9324
22.9886 | -35.3977 | -32.938
22.9948 | -35.3981 | -32.9437
23.0009 | -35.3985 | -32.9493
23.007 | -35.399 | -32.9546
23.0131 | -35.3994 | -32.9599
23.0192 | -35.3998 | -32.9647
23.0253 | -35.4002 | -32.9705
23.0314 | -35.4006 | -32.9763
23.0375 | -35.4009 | -32.9823
23.0436 | -35.4011 | -32.9882
23.0497 | -35.4015 | -32.9944
23.0558 | -35.402 | -33.0005
23.0619 | -35.4024 | -33.0057
23.068 | -35.4027 | -33.0106
23.0741 | -35.4031 | -33.0166
23.0802 | -35.4035 | -33.0216
23.0863 | -35.4039 | -33.0271
23.0924 | -35.4042 | -33.0331
23.0985 | -35.4045 | -33.039
23.1046 | -35.4049 | -33.0449
23.1107 | -35.4053 | -33.051
23.1168 | -35.4056 | -33.0569
23.1229 | -35.406 | -33.063
23.129 | -35.4064 | -33.0686
23.1351 | -35.4068 | -33.0746
23.1413 | -35.4072 | -33.0801
23.1474 | -35.4076 | -33.086
23.1535 | -35.408 | -33.0919
23.1596 | -35.4083 | -33.0973
23.1657 | -35.4087 | -33.1028
23.1718 | -35.4091 | -33.1085
23.1779 | -35.4094 | -33.1144
23.184 | -35.4098 | -33.1195
23.1901 | -35.4102 | -33.1257
23.1962 | -35.4106 | -33.1315
23.2023 | -35.411 | -33.137
23.2084 | -35.4114 | -33.1427
23.2145 | -35.4117 | -33.1481
23.2206 | -35.4122 | -33.1535
23.2267 | -35.4125 | -33.1591
23.2328 | -35.4129 | -33.1649
23.2389 | -35.4133 | -33.1706
23.245 | -35.4137 | -33.1763
23.2511 | -35.4141 | -33.1814
23.2572 | -35.4145 | -33.1867
23.2633 | -35.4149 | -33.1923
23.2694 | -35.4153 | -33.1979
23.2755 | -35.4157 | -33.204
23.2817 | -35.4161 | -33.2096
23.2878 | -35.4165 | -33.2151
23.2939 | -35.4169 | -33.2211
23.3 | -35.4173 | -33.2265
23.3061 | -35.4178 | -33.2318
23.3122 | -35.4181 | -33.2378
23.3183 | -35.4184 | -33.2435
23.3244 | -35.4188 | -33.249
23.3305 | -35.4192 | -33.2542
23.3366 | -35.4196 | -33.2602
23.3427 | -35.42 | -33.2655
23.3488 | -35.4203 | -33.2708
23.3549 | -35.4207 | -33.277
23.361 | -35.4212 | -33.2831
23.3671 | -35.4217 | -33.289
23.3732 | -35.422 | -33.2944
23.3793 | -35.4225 | -33.3001
23.3854 | -35.4229 | -33.3056
23.3915 | -35.4233 | -33.3104
23.3976 | -35.4238 | -33.3165
23.4037 | -35.4241 | -33.3221
23.4098 | -35.4245 | -33.3278
23.4159 | -35.4249 | -33.3336
23.422 | -35.4252 | -33.3396
23.4282 | -35.4256 | -33.3456
23.4343 | -35.426 | -33.3514
23.4404 | -35.4264 | -33.3574
23.4465 | -35.4266 | -33.3627
23.4526 | -35.4269 | -33.3688
23.4587 | -35.4273 | -33.3748
23.4648 | -35.4275 | -33.3802
23.4709 | -35.4279 | -33.3855
23.477 | -35.4282 | -33.3916
23.4831 | -35.4286 | -33.3967
23.4892 | -35.4289 | -33.4027
23.4953 | -35.4293 | -33.4085
23.5014 | -35.4296 | -33.4144
23.5075 | -35.43 | -33.4201
23.5136 | -35.4304 | -33.4255
23.5197 | -35.4307 | -33.4315
23.5258 | -35.4311 | -33.4374
23.5319 | -35.4315 | -33.4425
23.538 | -35.4319 | -33.4477
23.5441 | -35.4322 | -33.4534
23.5502 | -35.4325 | -33.4595
23.5563 | -35.4329 | -33.465
23.5624 | -35.4333 | -33.4704
23.5686 | -35.4336 | -33.4764
23.5747 | -35.434 | -33.4819
23.5808 | -35.4344 | -33.4875
23.5869 | -35.4348 | -33.4929
23.593 | -35.4352 | -33.4983
23.5991 | -35.4356 | -33.5036
23.6052 | -35.4359 | -33.509
23.6113 | -35.4363 | -33.5143
23.6174 | -35.4366 | -33.5198
23.6235 | -35.437 | -33.5252
23.6296 | -35.4373 | -33.5313
23.6357 | -35.4377 | -33.5369
23.6418 | -35.4381 | -33.5427
23.6479 | -35.4385 | -33.5481
23.654 | -35.4389 | -33.5544
23.6601 | -35.4392 | -33.5598
23.6662 | -35.4396 | -33.5655
23.6723 | -35.44 | -33.5702
23.6784 | -35.4404 | -33.576
23.6845 | -35.4409 | -33.5817
23.6906 | -35.4412 | -33.587
23.6967 | -35.4415 | -33.5933
23.7028 | -35.4419 | -33.5994
23.7089 | -35.4421 | -33.6062
23.7151 | -35.4425 | -33.6116
23.7212 | -35.4429 | -33.6184
23.7273 | -35.4431 | -33.6238
23.7334 | -35.4435 | -33.6304
23.7395 | -35.4438 | -33.6353
23.7456 | -35.4442 | -33.6417
23.7517 | -35.4447 | -33.6482
23.7578 | -35.4449 | -33.6537
23.7639 | -35.4455 | -33.6593
23.77 | -35.4459 | -33.6649
23.7761 | -35.4463 | -33.6708
23.7822 | -35.4467 | -33.6762
23.7883 | -35.447 | -33.6815
23.7944 | -35.4474 | -33.6871
23.8005 | -35.4478 | -33.6923
23.8066 | -35.4483 | -33.6979
23.8127 | -35.4486 | -33.7034
23.8188 | -35.449 | -33.7085
23.8249 | -35.4493 | -33.7143
23.831 | -35.4497 | -33.7202
23.8371 | -35.4501 | -33.7257
23.8432 | -35.4505 | -33.7312
23.8493 | -35.4508 | -33.7367
23.8555 | -35.4512 | -33.7423
23.8616 | -35.4516 | -33.7478
23.8677 | -35.452 | -33.7531
23.8738 | -35.4523 | -33.7588
23.8799 | -35.4527 | -33.7646
23.886 | -35.4531 | -33.7701
23.8921 | -35.4535 | -33.7763
23.8982 | -35.4539 | -33.7817
23.9043 | -35.4542 | -33.7875
23.9104 | -35.4545 | -33.7935
23.9165 | -35.4549 | -33.7985
23.9226 | -35.4552 | -33.8038
23.9287 | -35.4555 | -33.8092
23.9348 | -35.4559 | -33.8149
23.9409 | -35.4563 | -33.8209
23.947 | -35.4567 | -33.8263
23.9531 | -35.4571 | -33.8312
23.9592 | -35.4574 | -33.8366
23.9653 | -35.4577 | -33.8419
23.9714 | -35.458 | -33.8474
23.9775 | -35.4584 | -33.8531
23.9836 | -35.4588 | -33.8581
23.9897 | -35.4593 | -33.8642
23.9958 | -35.4596 | -33.8693
24.002 | -35.46 | -33.8748
24.0081 | -35.4603 | -33.8803
24.0142 | -35.4607 | -33.886
24.0203 | -35.461 | -33.8919
24.0264 | -35.4614 | -33.898
24.0325 | -35.4618 | -33.9036
24.0386 | -35.4621 | -33.9095
24.0447 | -35.4625 | -33.9149
24.0508 | -35.4628 | -33.9207
24.0569 | -35.4632 | -33.9263
24.063 | -35.4636 | -33.9319
24.0691 | -35.4639 | -33.937
24.0752 | -35.4642 | -33.9428
24.0813 | -35.4646 | -33.9489
24.0874 | -35.465 | -33.9544
24.0935 | -35.4653 | -33.9599
24.0996 | -35.4655 | -33.9654
24.1057 | -35.4659 | -33.9706
24.1118 | -35.4662 | -33.9756
24.1179 | -35.4666 | -33.9812
24.124 | -35.467 | -33.9863
24.1301 | -35.4673 | -33.9913
24.1362 | -35.4676 | -33.9973
24.1424 | -35.468 | -34.0033
24.1485 | -35.4684 | -34.0086
24.1546 | -35.4689 | -34.0141
24.1607 | -35.4692 | -34.0193
24.1668 | -35.4695 | -34.0246
24.1729 | -35.4698 | -34.0295
24.179 | -35.4701 | -34.0348
24.1851 | -35.4705 | -34.0409
24.1912 | -35.4708 | -34.0471
24.1973 | -35.4711 | -34.0534
24.2034 | -35.4715 | -34.0603
24.2095 | -35.4718 | -34.0654
24.2156 | -35.4722 | -34.0707
24.2217 | -35.4725 | -34.0763
24.2278 | -35.4728 | -34.0825
24.2339 | -35.4733 | -34.0878
24.24 | -35.4736 | -34.0927
24.2461 | -35.474 | -34.0991
24.2522 | -35.4744 | -34.1043
24.2583 | -35.4748 | -34.1105
24.2644 | -35.4753 | -34.1164
24.2705 | -35.4757 | -34.1235
24.2766 | -35.4761 | -34.1292
24.2827 | -35.4765 | -34.1347
24.2889 | -35.4768 | -34.14
24.295 | -35.4772 | -34.1461
24.3011 | -35.4774 | -34.1517
24.3072 | -35.4779 | -34.1578
24.3133 | -35.4782 | -34.1633
24.3194 | -35.4785 | -34.1687
24.3255 | -35.4789 | -34.1738
24.3316 | -35.4792 | -34.1798
24.3377 | -35.4795 | -34.1853
24.3438 | -35.4797 | -34.1915
24.3499 | -35.48 | -34.1967
24.356 | -35.4803 | -34.2021
24.3621 | -35.4807 | -34.2077
24.3682 | -35.481 | -34.2137
24.3743 | -35.4814 | -34.2192
24.3804 | -35.4816 | -34.225
24.3865 | -35.4819 | -34.2307
24.3926 | -35.4822 | -34.2366
24.3987 | -35.4824 | -34.2424
24.4048 | -35.4829 | -34.2485
24.4109 | -35.4834 | -34.2543
24.417 | -35.4837 | -34.2596
24.4231 | -35.484 | -34.2653
24.4293 | -35.4844 | -34.2708
24.4354 | -35.4847 | -34.2765
24.4415 | -35.4852 | -34.2823
24.4476 | -35.4855 | -34.2875
24.4537 | -35.4859 | -34.2931
24.4598 | -35.4862 | -34.2987
24.4659 | -35.4866 | -34.304
24.472 | -35.4869 | -34.3101
24.4781 | -35.4873 | -34.3161
24.4842 | -35.4877 | -34.3226
24.4903 | -35.4879 | -34.328
24.4964 | -35.4884 | -34.334
24.5025 | -35.4888 | -34.3398
24.5086 | -35.4891 | -34.3454
24.5147 | -35.4896 | -34.3508
24.5208 | -35.4901 | -34.3557
24.5269 | -35.4905 | -34.3612
24.533 | -35.4909 | -34.3667
24.5391 | -35.4913 | -34.3724
24.5452 | -35.4916 | -34.3779
24.5513 | -35.492 | -34.3836
24.5574 | -35.4925 | -34.3894
24.5635 | -35.4929 | -34.395
24.5696 | -35.4933 | -34.4002
24.5758 | -35.4937 | -34.4056
24.5819 | -35.494 | -34.4108
24.588 | -35.4945 | -34.4163
24.5941 | -35.4949 | -34.4218
24.6002 | -35.4952 | -34.4271
24.6063 | -35.4955 | -34.4322
24.6124 | -35.4959 | -34.4375
24.6185 | -35.4962 | -34.4437
24.6246 | -35.4967 | -34.4497
24.6307 | -35.4971 | -34.4552
24.6368 | -35.4975 | -34.4607
24.6429 | -35.4978 | -34.4654
24.649 | -35.4982 | -34.4714
24.6551 | -35.4985 | -34.4773
24.6612 | -35.499 | -34.4833
24.6673 | -35.4992 | -34.4886
24.6734 | -35.4997 | -34.4946
24.6795 | -35.5 | -34.5005
24.6856 | -35.5003 | -34.506
24.6917 | -35.5007 | -34.5116
24.6978 | -35.5012 | -34.5167
24.7039 | -35.5016 | -34.5225
24.71 | -35.5021 | -34.5285
24.7162 | -35.5025 | -34.5345
24.7223 | -35.5028 | -34.5404
24.7284 | -35.5033 | -34.5456
24.7345 | -35.5037 | -34.5512
24.7406 | -35.5039 | -34.557
24.7467 | -35.5043 | -34.5631
24.7528 | -35.5048 | -34.5686
24.7589 | -35.5051 | -34.5744
24.765 | -35.5055 | -34.5796
24.7711 | -35.5058 | -34.5856
24.7772 | -35.5061 | -34.5909
24.7833 | -35.5065 | -34.596
24.7894 | -35.5069 | -34.6015
24.7955 | -35.5073 | -34.6073
24.8016 | -35.5076 | -34.613
24.8077 | -35.508 | -34.619
24.8138 | -35.5084 | -34.6246
24.8199 | -35.5087 | -34.6305
24.826 | -35.509 | -34.6362
24.8321 | -35.5094 | -34.6421
24.8382 | -35.5098 | -34.6479
24.8443 | -35.5101 | -34.654
24.8504 | -35.5105 | -34.6597
24.8565 | -35.5109 | -34.6656
24.8627 | -35.5113 | -34.6715
24.8688 | -35.5117 | -34.6769
24.8749 | -35.5121 | -34.6826
24.881 | -35.5124 | -34.6887
24.8871 | -35.5128 | -34.6945
24.8932 | -35.5133 | -34.6997
24.8993 | -35.5136 | -34.705
24.9054 | -35.514 | -34.7108
24.9115 | -35.5143 | -34.7165
24.9176 | -35.5147 | -34.7221
24.9237 | -35.5151 | -34.7277
24.9298 | -35.5154 | -34.7336
24.9359 | -35.5158 | -34.7394
24.942 | -35.5161 | -34.7453
24.9481 | -35.5165 | -34.7509
24.9542 | -35.5169 | -34.7568
24.9603 | -35.5172 | -34.7626
24.9664 | -35.5176 | -34.7683
24.9725 | -35.518 | -34.7737
24.9786 | -35.5183 | -34.7797
24.9847 | -35.5187 | -34.785
24.9908 | -35.5192 | -34.7909
24.9969 | -35.5195 | -34.797
25.0031 | -35.52 | -34.8031
25.0092 | -35.5204 | -34.8086
25.0153 | -35.5207 | -34.8144
25.0214 | -35.5211 | -34.8199
25.0275 | -35.5215 | -34.8252
25.0336 | -35.5219 | -34.8311
25.0397 | -35.5223 | -34.8367
25.0458 | -35.5227 | -34.8432
25.0519 | -35.523 | -34.849
25.058 | -35.5234 | -34.8551
25.0641 | -35.5237 | -34.8604
25.0702 | -35.5241 | -34.8663
25.0763 | -35.5244 | -34.872
25.0824 | -35.5249 | -34.8774
25.0885 | -35.5251 | -34.8829
25.0946 | -35.5255 | -34.8886
25.1007 | -35.5258 | -34.8946
25.1068 | -35.5261 | -34.9002
25.1129 | -35.5264 | -34.9064
25.119 | -35.5267 | -34.912
25.1251 | -35.527 | -34.9176
25.1312 | -35.5275 | -34.9234
25.1373 | -35.5278 | -34.9289
25.1435 | -35.5283 | -34.9345
25.1496 | -35.5288 | -34.9399
25.1557 | -35.5291 | -34.9456
25.1618 | -35.5295 | -34.9513
25.1679 | -35.5299 | -34.9568
25.174 | -35.5303 | -34.9627
25.1801 | -35.5307 | -34.9694
25.1862 | -35.531 | -34.9755
25.1923 | -35.5314 | -34.9812
25.1984 | -35.5318 | -34.9871
25.2045 | -35.5322 | -34.9935
25.2106 | -35.5326 | -34.9996
25.2167 | -35.533 | -35.0058
25.2228 | -35.5333 | -35.0118
25.2289 | -35.5337 | -35.0177
25.235 | -35.534 | -35.0233
25.2411 | -35.5343 | -35.0288
25.2472 | -35.5347 | -35.0342
25.2533 | -35.5351 | -35.0401
25.2594 | -35.5354 | -35.046
25.2655 | -35.5358 | -35.052
25.2716 | -35.5361 | -35.0576
25.2777 | -35.5364 | -35.064
25.2838 | -35.5367 | -35.0694
25.29 | -35.5371 | -35.0751
25.2961 | -35.5374 | -35.0808
25.3022 | -35.5378 | -35.0857
25.3083 | -35.5382 | -35.0914
25.3144 | -35.5386 | -35.0968
25.3205 | -35.5389 | -35.1022
25.3266 | -35.5393 | -35.1077
25.3327 | -35.5397 | -35.1137
25.3388 | -35.54 | -35.1196
25.3449 | -35.5404 | -35.1254
25.351 | -35.5407 | -35.131
25.3571 | -35.5411 | -35.1366
25.3632 | -35.5414 | -35.1417
25.3693 | -35.5417 | -35.1477
25.3754 | -35.542 | -35.1535
25.3815 | -35.5425 | -35.1591
25.3876 | -35.5428 | -35.1652
25.3937 | -35.5431 | -35.1711
25.3998 | -35.5435 | -35.1765
25.4059 | -35.5439 | -35.1823
25.412 | -35.5442 | -35.1887
25.4181 | -35.5445 | -35.195
25.4242 | -35.5449 | -35.2011
25.4304 | -35.5452 | -35.2066
25.4365 | -35.5456 | -35.2128
25.4426 | -35.546 | -35.2184
25.4487 | -35.5463 | -35.2239
25.4548 | -35.5467 | -35.23
25.4609 | -35.5471 | -35.2356
25.467 | -35.5475 | -35.2416
25.4731 | -35.5478 | -35.2476
25.4792 | -35.5482 | -35.2532
25.4853 | -35.5486 | -35.2582
25.4914 | -35.5489 | -35.2641
25.4975 | -35.5493 | -35.2704
25.5036 | -35.5497 | -35.277
25.5097 | -35.5501 | -35.283
25.5158 | -35.5504 | -35.289
25.5219 | -35.5507 | -35.2945
25.528 | -35.551 | -35.3002
25.5341 | -35.5514 | -35.3058
25.5402 | -35.5518 | -35.312
25.5463 | -35.5522 | -35.3176
25.5524 | -35.5526 | -35.3234
25.5585 | -35.5529 | -35.3291
25.5646 | -35.5533 | -35.3352
25.5707 | -35.5537 | -35.3413
25.5769 | -35.554 | -35.3477
25.583 | -35.5545 | -35.3537
25.5891 | -35.5549 | -35.3594
25.5952 | -35.5553 | -35.3646
25.6013 | -35.5556 | -35.37
25.6074 | -35.556 | -35.3758
25.6135 | -35.5564 | -35.3815
25.6196 | -35.5567 | -35.3879
25.6257 | -35.5571 | -35.3941
25.6318 | -35.5575 | -35.3997
25.6379 | -35.5578 | -35.4053
25.644 | -35.5582 | -35.4114
25.6501 | -35.5586 | -35.4165
25.6562 | -35.559 | -35.4223
25.6623 | -35.5593 | -35.4279
25.6684 | -35.5597 | -35.4337
25.6745 | -35.56 | -35.439
25.6806 | -35.5603 | -35.4444
25.6867 | -35.5607 | -35.4509
25.6928 | -35.561 | -35.4565
25.6989 | -35.5613 | -35.4618
25.705 | -35.5616 | -35.4679
25.7111 | -35.5619 | -35.4736
25.7173 | -35.5623 | -35.4793
25.7234 | -35.5627 | -35.4855
25.7295 | -35.563 | -35.4913
25.7356 | -35.5634 | -35.4972
25.7417 | -35.5638 | -35.5026
25.7478 | -35.5642 | -35.5081
25.7539 | -35.5646 | -35.5131
25.76 | -35.5649 | -35.5187
25.7661 | -35.5652 | -35.5246
25.7722 | -35.5655 | -35.5303
25.7783 | -35.5659 | -35.5357
25.7844 | -35.5662 | -35.5419
25.7905 | -35.5664 | -35.5479
25.7966 | -35.5668 | -35.5538
25.8027 | -35.5671 | -35.5596
25.8088 | -35.5675 | -35.5658
25.8149 | -35.5679 | -35.5717
25.821 | -35.5683 | -35.5779
25.8271 | -35.5687 | -35.5842
25.8332 | -35.5689 | -35.5898
25.8393 | -35.5693 | -35.595
25.8454 | -35.5697 | -35.6004
25.8515 | -35.5701 | -35.6067
25.8576 | -35.5704 | -35.6124
25.8638 | -35.5707 | -35.6182
25.8699 | -35.5711 | -35.6238
25.876 | -35.5715 | -35.6297
25.8821 | -35.5719 | -35.6355
25.8882 | -35.5721 | -35.6409
25.8943 | -35.5726 | -35.6468
25.9004 | -35.5729 | -35.6521
25.9065 | -35.5732 | -35.6581
25.9126 | -35.5736 | -35.6636
25.9187 | -35.574 | -35.6693
25.9248 | -35.5743 | -35.6748
25.9309 | -35.5747 | -35.6805
25.937 | -35.5751 | -35.6858
25.9431 | -35.5755 | -35.6918
25.9492 | -35.5759 | -35.6971
25.9553 | -35.5762 | -35.7024
25.9614 | -35.5766 | -35.7081
25.9675 | -35.577 | -35.7138
25.9736 | -35.5774 | -35.7197
25.9797 | -35.5777 | -35.7258
25.9858 | -35.5781 | -35.7318
25.9919 | -35.5785 | -35.7376
25.998 | -35.5789 | -35.7431
26.0042 | -35.5793 | -35.7491
26.0103 | -35.5797 | -35.7544
26.0164 | -35.58 | -35.7598
26.0225 | -35.5803 | -35.7663
26.0286 | -35.5806 | -35.772
26.0347 | -35.5811 | -35.7782
26.0408 | -35.5815 | -35.7842
26.0469 | -35.5819 | -35.7894
26.053 | -35.5822 | -35.7951
26.0591 | -35.5826 | -35.8004
26.0652 | -35.583 | -35.8067
26.0713 | -35.5834 | -35.8126
26.0774 | -35.5837 | -35.8181
26.0835 | -35.584 | -35.8239
26.0896 | -35.5843 | -35.8293
26.0957 | -35.5847 | -35.8347
26.1018 | -35.5851 | -35.8406
26.1079 | -35.5855 | -35.8463
26.114 | -35.5858 | -35.8522
26.1201 | -35.5862 | -35.8579
26.1262 | -35.5865 | -35.8638
26.1323 | -35.5869 | -35.8695
26.1384 | -35.5873 | -35.8749
26.1445 | -35.5876 | -35.8808
26.1507 | -35.588 | -35.8867
26.1568 | -35.5883 | -35.893
26.1629 | -35.5887 | -35.8986
26.169 | -35.589 | -35.9046
26.1751 | -35.5894 | -35.911
26.1812 | -35.5898 | -35.9166
26.1873 | -35.5901 | -35.9221
26.1934 | -35.5905 | -35.9275
26.1995 | -35.5909 | -35.9335
26.2056 | -35.5913 | -35.9386
26.2117 | -35.5917 | -35.9451
26.2178 | -35.592 | -35.9512
26.2239 | -35.5923 | -35.9563
26.23 | -35.5927 | -35.9626
26.2361 | -35.593 | -35.9683
26.2422 | -35.5933 | -35.9743
26.2483 | -35.5937 | -35.9804
26.2544 | -35.594 | -35.9865
26.2605 | -35.5944 | -35.9923
26.2666 | -35.5947 | -35.9979
26.2727 | -35.595 | -36.0038
26.2788 | -35.5954 | -36.0088
26.2849 | -35.5958 | -36.0144
26.2911 | -35.5962 | -36.02
26.2972 | -35.5966 | -36.0258
26.3033 | -35.597 | -36.0319
26.3094 | -35.5973 | -36.0378
26.3155 | -35.5977 | -36.0427
26.3216 | -35.598 | -36.0484
26.3277 | -35.5984 | -36.0539
26.3338 | -35.5987 | -36.0592
26.3399 | -35.5991 | -36.0647
26.346 | -35.5994 | -36.07
26.3521 | -35.5997 | -36.076
26.3582 | -35.6 | -36.0822
26.3643 | -35.6003 | -36.0879
26.3704 | -35.6007 | -36.0937
26.3765 | -35.6011 | -36.0993
26.3826 | -35.6014 | -36.1046
26.3887 | -35.6018 | -36.1104
26.3948 | -35.6022 | -36.1159
26.4009 | -35.6025 | -36.1215
26.407 | -35.6029 | -36.1271
26.4131 | -35.6033 | -36.1331
26.4192 | -35.6038 | -36.1382
26.4253 | -35.6041 | -36.1435
26.4314 | -35.6045 | -36.1492
26.4376 | -35.6048 | -36.1548
26.4437 | -35.6052 | -36.161
26.4498 | -35.6056 | -36.167
26.4559 | -35.6059 | -36.172
26.462 | -35.6063 | -36.1776
26.4681 | -35.6066 | -36.1836
26.4742 | -35.6069 | -36.1888
26.4803 | -35.6072 | -36.195
26.4864 | -35.6077 | -36.201
26.4925 | -35.608 | -36.2068
26.4986 | -35.6083 | -36.2127
26.5047 | -35.6086 | -36.2186
26.5108 | -35.609 | -36.2246
26.5169 | -35.6094 | -36.2303
26.523 | -35.6097 | -36.236
26.5291 | -35.6101 | -36.2413
26.5352 | -35.6103 | -36.2469
26.5413 | -35.6107 | -36.2529
26.5474 | -35.611 | -36.2584
26.5535 | -35.6114 | -36.2644
26.5596 | -35.6117 | -36.2695
26.5657 | -35.6121 | -36.2752
26.5718 | -35.6125 | -36.2807
26.578 | -35.6128 | -36.2862
26.5841 | -35.6132 | -36.2922
26.5902 | -35.6135 | -36.2976
26.5963 | -35.6138 | -36.3037
26.6024 | -35.6141 | -36.3092
26.6085 | -35.6144 | -36.3145
26.6146 | -35.6148 | -36.3205
26.6207 | -35.6152 | -36.3264
26.6268 | -35.6156 | -36.3327
26.6329 | -35.6159 | -36.3379
26.639 | -35.6162 | -36.3438
26.6451 | -35.6165 | -36.3495
26.6512 | -35.6169 | -36.355
26.6573 | -35.6172 | -36.3607
26.6634 | -35.6174 | -36.3655
26.6695 | -35.6177 | -36.371
26.6756 | -35.618 | -36.377
26.6817 | -35.6183 | -36.3826
26.6878 | -35.6187 | -36.3877
26.6939 | -35.619 | -36.3939
26.7 | -35.6193 | -36.3992
26.7061 | -35.6196 | -36.4049
26.7122 | -35.6199 | -36.4106
26.7183 | -35.6201 | -36.4166
26.7245 | -35.6205 | -36.4224
26.7306 | -35.6207 | -36.4279
26.7367 | -35.621 | -36.4328
26.7428 | -35.6214 | -36.4381
26.7489 | -35.6218 | -36.4442
26.755 | -35.6221 | -36.4499
26.7611 | -35.6225 | -36.4559
26.7672 | -35.6227 | -36.4621
26.7733 | -35.6231 | -36.4677
26.7794 | -35.6234 | -36.4733
26.7855 | -35.6238 | -36.4793
26.7916 | -35.6241 | -36.4854
26.7977 | -35.6245 | -36.4908
26.8038 | -35.6249 | -36.4958
26.8099 | -35.6254 | -36.5013
26.816 | -35.6257 | -36.5069
26.8221 | -35.626 | -36.5125
26.8282 | -35.6264 | -36.5184
26.8343 | -35.6267 | -36.5241
26.8404 | -35.6271 | -36.5293
26.8465 | -35.6275 | -36.5352
26.8526 | -35.628 | -36.5411
26.8587 | -35.6283 | -36.5465
26.8649 | -35.6287 | -36.5521
26.871 | -35.629 | -36.5579
26.8771 | -35.6293 | -36.5629
26.8832 | -35.6296 | -36.568
26.8893 | -35.6299 | -36.5737
26.8954 | -35.6303 | -36.5794
26.9015 | -35.6306 | -36.5852
26.9076 | -35.631 | -36.5908
26.9137 | -35.6313 | -36.5957
26.9198 | -35.6317 | -36.601
26.9259 | -35.632 | -36.6064
26.932 | -35.6324 | -36.6113
26.9381 | -35.6328 | -36.617
26.9442 | -35.6331 | -36.6219
26.9503 | -35.6334 | -36.628
26.9564 | -35.6338 | -36.6338
26.9625 | -35.634 | -36.6392
26.9686 | -35.6344 | -36.6451
26.9747 | -35.6348 | -36.6508
26.9808 | -35.6351 | -36.6565
26.9869 | -35.6354 | -36.6623
26.993 | -35.6357 | -36.6677
26.9991 | -35.636 | -36.6733
27.0052 | -35.6363 | -36.6787
27.0114 | -35.6367 | -36.685
27.0175 | -35.637 | -36.6901
27.0236 | -35.6373 | -36.696
27.0297 | -35.6376 | -36.7018
27.0358 | -35.638 | -36.7077
27.0419 | -35.6384 | -36.7131
27.048 | -35.6387 | -36.7186
27.0541 | -35.639 | -36.724
27.0602 | -35.6393 | -36.7293
27.0663 | -35.6396 | -36.7347
27.0724 | -35.64 | -36.7404
27.0785 | -35.6403 | -36.7459
27.0846 | -35.6405 | -36.7517
27.0907 | -35.6408 | -36.7572
27.0968 | -35.6411 | -36.7628
27.1029 | -35.6415 | -36.7684
27.109 | -35.6418 | -36.7745
27.1151 | -35.6421 | -36.7805
27.1212 | -35.6426 | -36.7857
27.1273 | -35.643 | -36.7912
27.1334 | -35.6433 | -36.7965
27.1395 | -35.6436 | -36.8022
27.1456 | -35.6439 | -36.8082
27.1518 | -35.6442 | -36.814
27.1579 | -35.6446 | -36.8189
27.164 | -35.6449 | -36.824
27.1701 | -35.6453 | -36.8295
27.1762 | -35.6456 | -36.8348
27.1823 | -35.646 | -36.8404
27.1884 | -35.6463 | -36.8461
27.1945 | -35.6465 | -36.8516
27.2006 | -35.6468 | -36.8573
27.2067 | -35.6472 | -36.863
27.2128 | -35.6476 | -36.8686
27.2189 | -35.6479 | -36.8741
27.225 | -35.6481 | -36.8794
27.2311 | -35.6484 | -36.8848
27.2372 | -35.6487 | -36.8905
27.2433 | -35.6491 | -36.8958
27.2494 | -35.6494 | -36.9017
27.2555 | -35.6497 | -36.9069
27.2616 | -35.6499 | -36.9122
27.2677 | -35.6502 | -36.9175
27.2738 | -35.6505 | -36.9231
27.2799 | -35.6509 | -36.9285
27.286 | -35.6512 | -36.9341
27.2921 | -35.6514 | -36.9398
27.2983 | -35.6517 | -36.9451
27.3044 | -35.652 | -36.9506
27.3105 | -35.6524 | -36.9563
27.3166 | -35.6526 | -36.9616
27.3227 | -35.653 | -36.967
27.3288 | -35.6533 | -36.9723
27.3349 | -35.6535 | -36.9782
27.341 | -35.6538 | -36.9838
27.3471 | -35.6541 | -36.9891
27.3532 | -35.6544 | -36.9943
27.3593 | -35.6548 | -37.0001
27.3654 | -35.6552 | -37.0057
27.3715 | -35.6554 | -37.0108
27.3776 | -35.6557 | -37.0159
27.3837 | -35.656 | -37.0216
27.3898 | -35.6564 | -37.0269
27.3959 | -35.6568 | -37.0325
27.402 | -35.6571 | -37.0378
27.4081 | -35.6574 | -37.0436
27.4142 | -35.6578 | -37.0493
27.4203 | -35.6582 | -37.0542
27.4264 | -35.6585 | -37.0599
27.4325 | -35.6588 | -37.0652
27.4387 | -35.6591 | -37.0705
27.4448 | -35.6595 | -37.076
27.4509 | -35.6598 | -37.0818
27.457 | -35.6601 | -37.0871
27.4631 | -35.6605 | -37.0926
27.4692 | -35.6608 | -37.098
27.4753 | -35.6611 | -37.1033
27.4814 | -35.6614 | -37.1086
27.4875 | -35.6617 | -37.1146
27.4936 | -35.6621 | -37.12
27.4997 | -35.6623 | -37.1253
27.5058 | -35.6626 | -37.1307
27.5119 | -35.6629 | -37.1362
27.518 | -35.6634 | -37.1415
27.5241 | -35.6637 | -37.1469
27.5302 | -35.6641 | -37.152
27.5363 | -35.6644 | -37.1583
27.5424 | -35.6647 | -37.1637
27.5485 | -35.665 | -37.1698
27.5546 | -35.6653 | -37.1756
27.5607 | -35.6657 | -37.1813
27.5668 | -35.666 | -37.1865
27.5729 | -35.6663 | -37.1919
27.5791 | -35.6667 | -37.1969
27.5852 | -35.6669 | -37.202
27.5913 | -35.6672 | -37.2078
27.5974 | -35.6676 | -37.2131
27.6035 | -35.668 | -37.2187
27.6096 | -35.6682 | -37.2239
27.6157 | -35.6686 | -37.2292
27.6218 | -35.6689 | -37.2349
27.6279 | -35.6692 | -37.2404
27.634 | -35.6695 | -37.2458
27.6401 | -35.6697 | -37.2512
27.6462 | -35.67 | -37.2567
27.6523 | -35.6703 | -37.2621
27.6584 | -35.6705 | -37.2672
27.6645 | -35.6709 | -37.2727
27.6706 | -35.6712 | -37.278
27.6767 | -35.6715 | -37.2834
27.6828 | -35.6718 | -37.2886
27.6889 | -35.6721 | -37.2937
27.695 | -35.6724 | -37.2988
27.7011 | -35.6727 | -37.304
27.7072 | -35.6731 | -37.3097
27.7133 | -35.6735 | -37.3149
27.7194 | -35.6737 | -37.3197
27.7256 | -35.674 | -37.3251
27.7317 | -35.6743 | -37.3305
27.7378 | -35.6746 | -37.3364
27.7439 | -35.6749 | -37.3415
27.75 | -35.6752 | -37.347
27.7561 | -35.6755 | -37.3523
27.7622 | -35.6758 | -37.358
27.7683 | -35.676 | -37.3636
27.7744 | -35.6763 | -37.3691
27.7805 | -35.6766 | -37.3746
27.7866 | -35.677 | -37.38
27.7927 | -35.6772 | -37.3859
27.7988 | -35.6776 | -37.3911
27.8049 | -35.6779 | -37.3967
27.811 | -35.6782 | -37.4017
27.8171 | -35.6784 | -37.4071
27.8232 | -35.6788 | -37.4119
27.8293 | -35.679 | -37.4173
27.8354 | -35.6794 | -37.423
27.8415 | -35.6798 | -37.4285
27.8476 | -35.6801 | -37.4334
27.8537 | -35.6804 | -37.439
27.8598 | -35.6807 | -37.4443
27.866 | -35.6809 | -37.4498
27.8721 | -35.6812 | -37.4553
27.8782 | -35.6815 | -37.4602
27.8843 | -35.6818 | -37.4655
27.8904 | -35.682 | -37.4704
27.8965 | -35.6823 | -37.4757
27.9026 | -35.6825 | -37.4809
27.9087 | -35.6828 | -37.4856
27.9148 | -35.6832 | -37.4911
27.9209 | -35.6835 | -37.4964
27.927 | -35.6838 | -37.5018
27.9331 | -35.6841 | -37.5067
27.9392 | -35.6844 | -37.5125
27.9453 | -35.6847 | -37.5178
27.9514 | -35.685 | -37.523
27.9575 | -35.6854 | -37.5285
27.9636 | -35.6857 | -37.5335
27.9697 | -35.686 | -37.5383
27.9758 | -35.6863 | -37.543
27.9819 | -35.6867 | -37.5486
27.988 | -35.687 | -37.5541
27.9941 | -35.6872 | -37.5594
28.0002 | -35.6875 | -37.5643
28.0063 | -35.6877 | -37.5697
28.0125 | -35.6881 | -37.5748
28.0186 | -35.6884 | -37.5808
28.0247 | -35.6886 | -37.5854
28.0308 | -35.6889 | -37.5904
28.0369 | -35.6892 | -37.5959
28.043 | -35.6896 | -37.6018
28.0491 | -35.6898 | -37.6075
28.0552 | -35.69 | -37.6127
28.0613 | -35.6904 | -37.6184
28.0674 | -35.6907 | -37.6234
28.0735 | -35.6908 | -37.6289

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.5238 | -42.0639
21.4908 | -35.5561 | -34.647
16.4948 | -34.9424 | -27.3528
12.6602 | -34.5802 | -21.1513
9.7171 | -34.3674 | -16.2941
7.4582 | -34.2225 | -12.3601
5.7244 | -34.116 | -9.1726
4.3936 | -34.0585 | -6.5922
3.3722 | -34.0322 | -4.5626
2.5883 | -34.0209 | -3.0243
1.9866 | -34.011 | -1.7955
1.5248 | -34.0145 | -0.80745
1.1703 | -34.0279 | -0.16615
1 | -34.026 | 0.15021
0.89822 | -34.0202 | 0.41975
0.68941 | -34.0242 | 0.99555
0.52919 | -34.0218 | 1.4403
0.40612 | -34.0347 | 2.0411
0.31171 | -34.0187 | 2.4939
0.23928 | -34.0226 | 3.2111
0.18358 | -34.0322 | 3.62
0.14091 | -34.0414 | 4.5031
0.10819 | -34.0635 | 5.6523
0.083017 | -34.076 | 7.1225
0.063705 | -34.1296 | 8.9724
0.048923 | -34.1966 | 11.4414
0.037479 | -34.2897 | 14.6316
0.028801 | -34.4619 | 18.6643
0.022078 | -34.7689 | 24.12
0.017023 | -35.262 | 30.554
0.013018 | -36.069 | 38.2231
0.010014 | -37.1789 | 46.1416

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003628 | -1.4408
0.007256 | -25.4493
0.010884 | -34.0148
0.014512 | -41.72
0.018141 | -43.7771
0.021769 | -47.1828
0.025397 | -48.4616
0.029025 | -50.803
0.032653 | -52.3536
0.036281 | -53.0838
0.039909 | -53.5383
0.043537 | -55.1487
0.047166 | -56.4367
0.050794 | -56.5744
0.054422 | -57.3133
0.05805 | -58.1019
0.061678 | -58.2387
0.065306 | -59.0125
0.068934 | -59.6649
0.072562 | -60.3368
0.07619 | -60.3003
0.079819 | -62.374
0.083447 | -61.9771
0.087075 | -61.3994
0.090703 | -62.2522
0.094331 | -64.6205
0.097959 | -64.818
0.10159 | -64.8493
0.10522 | -64.135
0.10884 | -64.1583
0.11247 | -64.3497
0.1161 | -65.1909
0.11973 | -65.9383
0.12336 | -65.8094
0.12698 | -66.8971
0.13061 | -68.3122
0.13424 | -67.9273
0.13787 | -67.3745
0.1415 | -67.8345
0.14513 | -69.7711
0.14875 | -69.0425
0.15238 | -68.203
0.15601 | -68.4921
0.15964 | -69.0605
0.16326 | -69.7793
0.16689 | -70.1836
0.17052 | -69.9385
0.17415 | -70.4825
0.17778 | -71.4614
0.18141 | -70.969
0.18503 | -72.044
0.18866 | -72.0874
0.19229 | -72.9853
0.19592 | -72.5906
0.19955 | -71.726
0.20317 | -71.2198
0.2068 | -72.4505
0.21043 | -73.1663
0.21406 | -73.3031
0.21769 | -73.0994
0.22132 | -71.8623
0.22494 | -71.6456
0.22857 | -73.9903
0.2322 | -75.6839
0.23583 | -75.1724
0.23946 | -75.1702
0.24308 | -73.7078
0.24671 | -73.5666
0.25034 | -73.1925
0.25397 | -73.6374
0.2576 | -74.0737
0.26122 | -74.3466
0.26485 | -74.3804
0.26848 | -74.4496
0.27211 | -74.6096
0.27574 | -74.7524
0.27936 | -74.9839
0.28299 | -74.8748
0.28662 | -74.3749
0.29025 | -74.5857
0.29388 | -75.8706
0.29751 | -76.2456
0.30113 | -77.28
0.30476 | -76.1846
0.30839 | -75.6936
0.31202 | -75.9817
0.31565 | -75.6446
0.31927 | -74.8955
0.3229 | -76.2617
0.32653 | -77.3085
0.33016 | -77.1043
0.33379 | -76.6579
0.33742 | -76.9535
0.34104 | -77.3517
0.34467 | -77.2804
0.3483 | -77.7927
0.35193 | -77.9667
0.35556 | -77.7419
0.35918 | -77.0134
0.36281 | -78.2653
0.36644 | -78.9029
0.37007 | -77.0243
0.3737 | -77.3534
0.37732 | -78.5125
0.38095 | -77.5045
0.38458 | -77.4295
0.38821 | -78.7421
0.39184 | -79.5821
0.39547 | -78.5742
0.39909 | -78.2325
0.40272 | -78.5916
0.40635 | -78.3799
0.40998 | -78.6807
0.41361 | -78.5423
0.41723 | -77.9707
0.42086 | -78.8779
0.42449 | -78.9371
0.42812 | -78.9925
0.43175 | -78.174
0.43537 | -79.8362
0.439 | -78.9319
0.44263 | -78.8369
0.44626 | -79.5589
0.44989 | -79.311
0.45352 | -78.7147
0.45714 | -79.1597
0.46077 | -79.3662
0.4644 | -80.0935
0.46803 | -78.9795
0.47165 | -79.7059
0.47528 | -80.4771
0.47891 | -79.6684
0.48254 | -79.6394
0.48617 | -80.5227
0.4898 | -81.222
0.49342 | -80.9029
0.49705 | -80.8152
0.50068 | -80.9113
0.50431 | -79.8063
0.50794 | -80.6714
0.51157 | -80.838
0.51519 | -80.0822
0.51882 | -79.9986
0.52245 | -80.6863
0.52608 | -81.3618
0.5297 | -81.3952
0.53333 | -81.1464
0.53696 | -81.1738
0.54059 | -81.2833
0.54422 | -81.0749
0.54785 | -80.9218
0.55147 | -81.2587
0.5551 | -81.5768
0.55873 | -81.6353
0.56236 | -81.402
0.56599 | -81.3218
0.56961 | -81.0463
0.57324 | -80.945
0.57687 | -81.4758
0.5805 | -82.009
0.58413 | -82.1484
0.58776 | -82.3186
0.59138 | -82.4786
0.59501 | -82.5553
0.59864 | -82.7025
0.60227 | -82.8921
0.6059 | -82.7425
0.60952 | -82.3266
0.61315 | -82.1545
0.61678 | -82.5266
0.62041 | -83.1979
0.62404 | -83.7068
0.62766 | -83.4255
0.63129 | -82.8125
0.63492 | -82.3873
0.63855 | -82.2881
0.64218 | -82.3387
0.6458 | -82.6555
0.64943 | -83.0755
0.65306 | -83.1794
0.65669 | -83.0971
0.66032 | -82.9446
0.66395 | -82.8542
0.66757 | -83.127
0.6712 | -83.4118
0.67483 | -83.3635
0.67846 | -83.2182
0.68209 | -83.1539
0.68571 | -82.9874
0.68934 | -83.1065
0.69297 | -83.669
0.6966 | -83.8915
0.70023 | -83.6925
0.70386 | -83.6622
0.70748 | -83.8334
0.71111 | -83.6952
0.71474 | -83.7508
0.71837 | -84.0404
0.722 | -84.2346
0.72562 | -84.4862
0.72925 | -84.4618
0.73288 | -84.229
0.73651 | -83.8523
0.74014 | -83.6503
0.74376 | -84.0333
0.74739 | -84.6368
0.75102 | -84.8151
0.75465 | -84.7247
0.75828 | -84.6203
0.76191 | -84.5699
0.76553 | -84.6668
0.76916 | -85.0362
0.77279 | -85.0528
0.77642 | -84.8669
0.78004 | -84.7468
0.78367 | -84.6295
0.7873 | -84.8003
0.79093 | -84.7011
0.79456 | -84.1817
0.79819 | -84.5209
0.80181 | -85.4113
0.80544 | -85.5905
0.80907 | -85.2314
0.8127 | -85.3331
0.81633 | -85.4605
0.81995 | -85.3261
0.82358 | -85.3193
0.82721 | -85.6003
0.83084 | -86.1248
0.83447 | -86.1346
0.8381 | -85.7173
0.84172 | -85.4949
0.84535 | -85.5344
0.84898 | -85.2275
0.85261 | -84.8206
0.85624 | -84.9528
0.85986 | -85.3903
0.86349 | -85.6661
0.86712 | -85.8736
0.87075 | -85.9719
0.87438 | -85.833
0.87801 | -85.644
0.88163 | -85.9055
0.88526 | -86.2919
0.88889 | -86.1393
0.89252 | -86.3215
0.89614 | -86.7269
0.89977 | -86.8355
0.9034 | -86.9356
0.90703 | -86.7538
0.91066 | -86.3724
0.91429 | -86.1798
0.91791 | -85.8716
0.92154 | -85.3826
0.92517 | -85.34
0.9288 | -85.4269
0.93243 | -85.427
0.93605 | -86.0082
0.93968 | -86.735
0.94331 | -86.8462
0.94694 | -86.4214
0.95057 | -85.9597
0.9542 | -85.8946
0.95782 | -86.6364
0.96145 | -87.5267
0.96508 | -87.3252
0.96871 | -86.7953
0.97234 | -86.8789
0.97596 | -86.8497
0.97959 | -86.1871
0.98322 | -85.8066
0.98685 | -86.1106
0.99048 | -86.5776
0.9941 | -86.438
0.99773 | -86.0306
1.0014 | -86.0889
1.005 | -86.5673
1.0086 | -86.8639
1.0122 | -86.993
1.0159 | -87.2959
1.0195 | -87.8728
1.0231 | -88.1135
1.0268 | -87.8582
1.0304 | -87.6381
1.034 | -87.7281
1.0376 | -87.9773
1.0413 | -87.9177
1.0449 | -87.4479
1.0485 | -87.2336
1.0522 | -87.3161
1.0558 | -87.5887
1.0594 | -87.8469
1.063 | -88.0926
1.0667 | -88.3981
1.0703 | -88.6033
1.0739 | -88.4902
1.0776 | -88.2483
1.0812 | -88.1926
1.0848 | -88.2586
1.0884 | -88.1429
1.0921 | -87.796
1.0957 | -87.4232
1.0993 | -87.2387
1.1029 | -87.3449
1.1066 | -87.6531
1.1102 | -88.0431
1.1138 | -88.4216
1.1175 | -88.6503
1.1211 | -88.7263
1.1247 | -88.7185
1.1283 | -88.6497
1.132 | -88.626
1.1356 | -88.7101
1.1392 | -88.8402
1.1429 | -88.9461
1.1465 | -89.0166
1.1501 | -89.0352
1.1537 | -88.8856
1.1574 | -88.5693
1.161 | -88.2808
1.1646 | -88.1527
1.1683 | -88.2457
1.1719 | -88.5309
1.1755 | -88.8025
1.1791 | -88.9293
1.1828 | -88.9862
1.1864 | -88.9304
1.19 | -88.6383
1.1937 | -88.2771
1.1973 | -88.1412
1.2009 | -88.3347
1.2045 | -88.7423
1.2082 | -89.1478
1.2118 | -89.3864
1.2154 | -89.5061
1.219 | -89.6569
1.2227 | -89.8282
1.2263 | -89.8955
1.2299 | -89.8871
1.2336 | -89.907
1.2372 | -89.8253
1.2408 | -89.3906
1.2444 | -88.6962
1.2481 | -88.1153
1.2517 | -87.9378
1.2553 | -88.1729
1.259 | -88.5156
1.2626 | -88.6851
1.2662 | -88.7305
1.2698 | -88.756
1.2735 | -88.8671
1.2771 | -89.1437
1.2807 | -89.4153
1.2844 | -89.5326
1.288 | -89.5425
1.2916 | -89.5092
1.2952 | -89.6175
1.2989 | -89.9727
1.3025 | -90.2911
1.3061 | -90.2834
1.3098 | -90.0571
1.3134 | -89.8138
1.317 | -89.6265
1.3206 | -89.5743
1.3243 | -89.8285
1.3279 | -90.3757
1.3315 | -90.7589
1.3351 | -90.6125
1.3388 | -90.2058
1.3424 | -89.7714
1.346 | -89.4023
1.3497 | -89.2813
1.3533 | -89.4933
1.3569 | -89.8801
1.3605 | -90.1988
1.3642 | -90.2921
1.3678 | -90.1668
1.3714 | -90.052
1.3751 | -90.196
1.3787 | -90.4797
1.3823 | -90.6132
1.3859 | -90.6215
1.3896 | -90.7605
1.3932 | -91.0077
1.3968 | -91.0106
1.4005 | -90.5999
1.4041 | -90.0872
1.4077 | -89.8067
1.4113 | -89.772
1.415 | -89.8248
1.4186 | -89.9445
1.4222 | -90.1295
1.4259 | -90.1918
1.4295 | -90.0766
1.4331 | -90.0848
1.4367 | -90.3769
1.4404 | -90.7788
1.444 | -91.039
1.4476 | -91.0157
1.4512 | -90.854
1.4549 | -90.8255
1.4585 | -90.9405
1.4621 | -91.0486
1.4658 | -91.0752
1.4694 | -91.0366
1.473 | -90.9229
1.4766 | -90.8309
1.4803 | -91.0021
1.4839 | -91.3383
1.4875 | -91.4242
1.4912 | -91.2885
1.4948 | -91.2648
1.4984 | -91.403
1.502 | -91.518
1.5057 | -91.4342
1.5093 | -91.2165
1.5129 | -91.124
1.5166 | -91.3099
1.5202 | -91.6715
1.5238 | -91.8976
1.5274 | -91.8141
1.5311 | -91.6241
1.5347 | -91.4176
1.5383 | -91.1422
1.5419 | -90.9447
1.5456 | -90.9863
1.5492 | -91.2527
1.5528 | -91.6211
1.5565 | -91.9416
1.5601 | -92.1259
1.5637 | -92.1718
1.5673 | -91.9819
1.571 | -91.7876
1.5746 | -91.5995
1.5782 | -91.5339
1.5819 | -91.6409
1.5855 | -91.8248
1.5891 | -91.925
1.5927 | -91.8419
1.5964 | -91.6784
1.6 | -91.6048
1.6036 | -91.7437
1.6073 | -92.0341
1.6109 | -92.2854
1.6145 | -92.3273
1.6181 | -92.1375
1.6218 | -91.7927
1.6254 | -91.4705
1.629 | -91.3566
1.6327 | -91.509
1.6363 | -91.8066
1.6399 | -92.0384
1.6435 | -92.051
1.6472 | -91.8817
1.6508 | -91.6867
1.6544 | -91.6079
1.6581 | -91.6743
1.6617 | -91.7927
1.6653 | -91.8615
1.6689 | -91.8626
1.6726 | -91.85
1.6762 | -91.8355
1.6798 | -91.7978
1.6834 | -91.7799
1.6871 | -91.8631
1.6907 | -92.0456
1.6943 | -92.1921
1.698 | -92.1831
1.7016 | -92.0964
1.7052 | -92.1231
1.7088 | -92.3395
1.7125 | -92.6592
1.7161 | -92.9577
1.7197 | -93.1687
1.7234 | -93.2803
1.727 | -93.2698
1.7306 | -93.1707
1.7342 | -93.1232
1.7379 | -93.2051
1.7415 | -93.3141
1.7451 | -93.2634
1.7488 | -92.9864
1.7524 | -92.5534
1.756 | -92.1391
1.7596 | -91.8993
1.7633 | -91.8841
1.7669 | -92.0391
1.7705 | -92.2574
1.7741 | -92.4556
1.7778 | -92.5593
1.7814 | -92.5826
1.785 | -92.6213
1.7887 | -92.7138
1.7923 | -92.8592
1.7959 | -92.9916
1.7995 | -93.0563
1.8032 | -93.0594
1.8068 | -93.0483
1.8104 | -93.0659
1.8141 | -93.0846
1.8177 | -93.0742
1.8213 | -93.0082
1.8249 | -92.8634
1.8286 | -92.6746
1.8322 | -92.4785
1.8358 | -92.2796
1.8395 | -92.1159
1.8431 | -92.0817
1.8467 | -92.2525
1.8503 | -92.5817
1.854 | -92.9064
1.8576 | -93.0835
1.8612 | -93.0545
1.8649 | -92.835
1.8685 | -92.4577
1.8721 | -92.0412
1.8757 | -91.7712
1.8794 | -91.7671
1.883 | -91.9582
1.8866 | -92.1842
1.8902 | -92.3534
1.8939 | -92.5077
1.8975 | -92.6853
1.9011 | -92.8973
1.9048 | -93.1284
1.9084 | -93.3106
1.912 | -93.3809
1.9156 | -93.4365
1.9193 | -93.5937
1.9229 | -93.8604
1.9265 | -94.0971
1.9302 | -94.1191
1.9338 | -93.8733
1.9374 | -93.508
1.941 | -93.2085
1.9447 | -92.9855
1.9483 | -92.7948
1.9519 | -92.6678
1.9556 | -92.638
1.9592 | -92.6562
1.9628 | -92.6064
1.9664 | -92.4792
1.9701 | -92.4202
1.9737 | -92.5755
1.9773 | -92.9096
1.981 | -93.1986
1.9846 | -93.2874
1.9882 | -93.2791
1.9918 | -93.336
1.9955 | -93.4912
1.9991 | -93.652
2.0027 | -93.7221
2.0063 | -93.6697
2.01 | -93.5578
2.0136 | -93.4974
2.0172 | -93.5214
2.0209 | -93.5682
2.0245 | -93.5479
2.0281 | -93.4734
2.0317 | -93.4768
2.0354 | -93.6277
2.039 | -93.8368
2.0426 | -93.9442
2.0463 | -93.885
2.0499 | -93.7824
2.0535 | -93.7853
2.0571 | -93.9276
2.0608 | -94.086
2.0644 | -94.0856
2.068 | -93.9094
2.0717 | -93.6516
2.0753 | -93.4013
2.0789 | -93.2383
2.0825 | -93.2354
2.0862 | -93.4523
2.0898 | -93.881
2.0934 | -94.2725
2.0971 | -94.5839
2.1007 | -94.7381
2.1043 | -94.7275
2.1079 | -94.6324
2.1116 | -94.5318
2.1152 | -94.4915
2.1188 | -94.5237
2.1224 | -94.6059
2.1261 | -94.7089
2.1297 | -94.8153
2.1333 | -94.8737
2.137 | -94.8654
2.1406 | -94.7691
2.1442 | -94.6072
2.1478 | -94.4169
2.1515 | -94.2704
2.1551 | -94.2045
2.1587 | -94.2228
2.1624 | -94.3079
2.166 | -94.4336
2.1696 | -94.5538
2.1732 | -94.6248
2.1769 | -94.6067
2.1805 | -94.5239
2.1841 | -94.4036
2.1878 | -94.2792
2.1914 | -94.1785
2.195 | -94.1325
2.1986 | -94.1601
2.2023 | -94.2603
2.2059 | -94.3903
2.2095 | -94.4866
2.2132 | -94.511
2.2168 | -94.452
2.2204 | -94.3133
2.224 | -94.1322
2.2277 | -93.9851
2.2313 | -93.9405
2.2349 | -94.0263
2.2385 | -94.2112
2.2422 | -94.4246
2.2458 | -94.5981
2.2494 | -94.7115
2.2531 | -94.7952
2.2567 | -94.898
2.2603 | -95.0343
2.2639 | -95.1313
2.2676 | -95.1064
2.2712 | -94.9457
2.2748 | -94.7071
2.2785 | -94.4741
2.2821 | -94.2974
2.2857 | -94.211
2.2893 | -94.228
2.293 | -94.3461
2.2966 | -94.5644
2.3002 | -94.8255
2.3039 | -95.026
2.3075 | -95.0446
2.3111 | -94.8452
2.3147 | -94.5296
2.3184 | -94.2474
2.322 | -94.123
2.3256 | -94.1919
2.3293 | -94.3764
2.3329 | -94.5387
2.3365 | -94.5855
2.3401 | -94.533
2.3438 | -94.4683
2.3474 | -94.4589
2.351 | -94.5022
2.3546 | -94.5612
2.3583 | -94.5882
2.3619 | -94.5668
2.3655 | -94.5207
2.3692 | -94.4907
2.3728 | -94.4866
2.3764 | -94.4811
2.38 | -94.4705
2.3837 | -94.4606
2.3873 | -94.4766
2.3909 | -94.5202
2.3946 | -94.5741
2.3982 | -94.624
2.4018 | -94.6506
2.4054 | -94.6435
2.4091 | -94.5998
2.4127 | -94.5446
2.4163 | -94.5398
2.42 | -94.6261
2.4236 | -94.8182
2.4272 | -95.0905
2.4308 | -95.4017
2.4345 | -95.6687
2.4381 | -95.8118
2.4417 | -95.8032
2.4454 | -95.6788
2.449 | -95.5261
2.4526 | -95.393
2.4562 | -95.2701
2.4599 | -95.141
2.4635 | -95.0396
2.4671 | -95.0321
2.4707 | -95.1615
2.4744 | -95.3869
2.478 | -95.6021
2.4816 | -95.7302
2.4853 | -95.7586
2.4889 | -95.7053
2.4925 | -95.5976
2.4961 | -95.49
2.4998 | -95.4548
2.5034 | -95.5129
2.507 | -95.5749
2.5107 | -95.5325
2.5143 | -95.3504
2.5179 | -95.0902
2.5215 | -94.8485
2.5252 | -94.7377
2.5288 | -94.8189
2.5324 | -95.0915
2.5361 | -95.4496
2.5397 | -95.735
2.5433 | -95.8245
2.5469 | -95.7259
2.5506 | -95.5326
2.5542 | -95.3776
2.5578 | -95.3349
2.5615 | -95.4
2.5651 | -95.5035
2.5687 | -95.5635
2.5723 | -95.5525
2.576 | -95.5237
2.5796 | -95.5399
2.5832 | -95.6581
2.5868 | -95.9021
2.5905 | -96.222
2.5941 | -96.5013
2.5977 | -96.631
2.6014 | -96.5907
2.605 | -96.4653
2.6086 | -96.359
2.6122 | -96.3079
2.6159 | -96.2504
2.6195 | -96.1393
2.6231 | -95.9883
2.6268 | -95.8123
2.6304 | -95.6423
2.634 | -95.5186
2.6376 | -95.4818
2.6413 | -95.544
2.6449 | -95.6765
2.6485 | -95.82
2.6522 | -95.9308
2.6558 | -95.9914
2.6594 | -96.0182
2.663 | -96.0068
2.6667 | -95.9845
2.6703 | -95.9545
2.6739 | -95.9235
2.6776 | -95.8778
2.6812 | -95.8139
2.6848 | -95.7335
2.6884 | -95.66
2.6921 | -95.6143
2.6957 | -95.5918
2.6993 | -95.5719
2.7029 | -95.5218
2.7066 | -95.4403
2.7102 | -95.358
2.7138 | -95.2895
2.7175 | -95.2458
2.7211 | -95.2187
2.7247 | -95.2041
2.7283 | -95.1753
2.732 | -95.1529
2.7356 | -95.1469
2.7392 | -95.1621
2.7429 | -95.1919
2.7465 | -95.2142
2.7501 | -95.2411
2.7537 | -95.2899
2.7574 | -95.3813
2.761 | -95.5407
2.7646 | -95.7438
2.7683 | -95.9444
2.7719 | -96.089
2.7755 | -96.1817
2.7791 | -96.2176
2.7828 | -96.2214
2.7864 | -96.207
2.79 | -96.1847
2.7937 | -96.1586
2.7973 | -96.1188
2.8009 | -96.0856
2.8045 | -96.0625
2.8082 | -96.0304
2.8118 | -95.9872
2.8154 | -95.9433
2.819 | -95.9153
2.8227 | -95.9055
2.8263 | -95.9377
2.8299 | -96.0184
2.8336 | -96.123
2.8372 | -96.2106
2.8408 | -96.2588
2.8444 | -96.258
2.8481 | -96.2146
2.8517 | -96.1549
2.8553 | -96.1274
2.859 | -96.1529
2.8626 | -96.2052
2.8662 | -96.2559
2.8698 | -96.2901
2.8735 | -96.2953
2.8771 | -96.2593
2.8807 | -96.1894
2.8844 | -96.1158
2.888 | -96.0711
2.8916 | -96.0643
2.8952 | -96.0879
2.8989 | -96.1416
2.9025 | -96.2024
2.9061 | -96.262
2.9098 | -96.3331
2.9134 | -96.4276
2.917 | -96.5359
2.9206 | -96.643
2.9243 | -96.7317
2.9279 | -96.7728
2.9315 | -96.7544
2.9351 | -96.6782
2.9388 | -96.5755
2.9424 | -96.4848
2.946 | -96.4196
2.9497 | -96.3855
2.9533 | -96.3926
2.9569 | -96.4081
2.9605 | -96.4174
2.9642 | -96.4174
2.9678 | -96.4063
2.9714 | -96.3934
2.9751 | -96.3941
2.9787 | -96.3817
2.9823 | -96.3673
2.9859 | -96.3561
2.9896 | -96.3537
2.9932 | -96.3692
2.9968 | -96.3993
3.0005 | -96.4293
3.0041 | -96.4788
3.0077 | -96.5437
3.0113 | -96.6341
3.015 | -96.7486
3.0186 | -96.8662
3.0222 | -96.9746
3.0259 | -97.0808
3.0295 | -97.1469
3.0331 | -97.1651
3.0367 | -97.1355
3.0404 | -97.0812
3.044 | -97.0158
3.0476 | -96.9699
3.0512 | -96.949
3.0549 | -96.9665
3.0585 | -97.0171
3.0621 | -97.0786
3.0658 | -97.1266
3.0694 | -97.1403
3.073 | -97.1229
3.0766 | -97.0909
3.0803 | -97.0674
3.0839 | -97.0754
3.0875 | -97.1136
3.0912 | -97.1673
3.0948 | -97.2205
3.0984 | -97.2799
3.102 | -97.3373
3.1057 | -97.3782
3.1093 | -97.3933
3.1129 | -97.3715
3.1166 | -97.3192
3.1202 | -97.2688
3.1238 | -97.2455
3.1274 | -97.2427
3.1311 | -97.2547
3.1347 | -97.2737
3.1383 | -97.2306
3.142 | -97.1921
3.1456 | -97.1345
3.1492 | -97.0667
3.1528 | -96.9987
3.1565 | -96.9454
3.1601 | -96.9094
3.1637 | -96.8921
3.1673 | -96.8894
3.171 | -96.8983
3.1746 | -96.9205
3.1782 | -96.9492
3.1819 | -96.9693
3.1855 | -96.9684
3.1891 | -96.9683
3.1927 | -96.9691
3.1964 | -96.9783
3.2 | -97.0072
3.2036 | -97.0463
3.2073 | -97.0868
3.2109 | -97.1162
3.2145 | -97.1366
3.2181 | -97.1525
3.2218 | -97.1706
3.2254 | -97.1946
3.229 | -97.2141
3.2327 | -97.2329
3.2363 | -97.2472
3.2399 | -97.2637
3.2435 | -97.27
3.2472 | -97.2776
3.2508 | -97.2756
3.2544 | -97.2576
3.258 | -97.2293
3.2617 | -97.2038
3.2653 | -97.1887
3.2689 | -97.168
3.2726 | -97.128
3.2762 | -97.0676
3.2798 | -96.9984
3.2834 | -96.9399
3.2871 | -96.9118
3.2907 | -96.9378
3.2943 | -97.0191
3.298 | -97.148
3.3016 | -97.2876
3.3052 | -97.4241
3.3088 | -97.5308
3.3125 | -97.5942
3.3161 | -97.5974
3.3197 | -97.5616
3.3234 | -97.5246
3.327 | -97.5067
3.3306 | -97.5059
3.3342 | -97.5358
3.3379 | -97.5991
3.3415 | -97.6802
3.3451 | -97.7442
3.3488 | -97.7772
3.3524 | -97.7742
3.356 | -97.7297
3.3596 | -97.6413
3.3633 | -97.5336
3.3669 | -97.4359
3.3705 | -97.3674
3.3742 | -97.3316
3.3778 | -97.3254
3.3814 | -97.3406
3.385 | -97.3776
3.3887 | -97.4192
3.3923 | -97.4675
3.3959 | -97.52
3.3995 | -97.5914
3.4032 | -97.6765
3.4068 | -97.7739
3.4104 | -97.8794
3.4141 | -97.9706
3.4177 | -98.0293
3.4213 | -98.0505
3.4249 | -98.0443
3.4286 | -98.0186
3.4322 | -97.9951
3.4358 | -97.981
3.4395 | -97.9761
3.4431 | -97.9841
3.4467 | -97.9948
3.4503 | -98.0044
3.454 | -98.0046
3.4576 | -97.9936
3.4612 | -97.9702
3.4649 | -97.9472
3.4685 | -97.9135
3.4721 | -97.8563
3.4757 | -97.7826
3.4794 | -97.7011
3.483 | -97.6282
3.4866 | -97.5869
3.4902 | -97.5833
3.4939 | -97.6127
3.4975 | -97.6687
3.5011 | -97.734
3.5048 | -97.7995
3.5084 | -97.8611
3.512 | -97.9065
3.5156 | -97.9158
3.5193 | -97.8793
3.5229 | -97.8275
3.5265 | -97.7908
3.5302 | -97.7951
3.5338 | -97.8644
3.5374 | -98.0038
3.541 | -98.1837
3.5447 | -98.381
3.5483 | -98.5711
3.5519 | -98.7054
3.5556 | -98.7629
3.5592 | -98.7397
3.5628 | -98.6445
3.5664 | -98.5171
3.5701 | -98.4111
3.5737 | -98.3493
3.5773 | -98.3267
3.581 | -98.3394
3.5846 | -98.3768
3.5882 | -98.4301
3.5918 | -98.496
3.5955 | -98.5637
3.5991 | -98.6256
3.6027 | -98.6641
3.6063 | -98.6869
3.61 | -98.674
3.6136 | -98.6306
3.6172 | -98.5681
3.6209 | -98.4899
3.6245 | -98.4066
3.6281 | -98.3377
3.6317 | -98.2928
3.6354 | -98.2712
3.639 | -98.25
3.6426 | -98.2058
3.6463 | -98.1268
3.6499 | -98.0412
3.6535 | -97.976
3.6571 | -97.9505
3.6608 | -98.0267
3.6644 | -98.0805
3.668 | -98.1661
3.6717 | -98.2584
3.6753 | -98.3506
3.6789 | -98.4376
3.6825 | -98.4961
3.6862 | -98.5283
3.6898 | -98.5277
3.6934 | -98.4778
3.6971 | -98.3999
3.7007 | -98.3065
3.7043 | -98.2134
3.7079 | -98.1184
3.7116 | -98.0243
3.7152 | -97.9514
3.7188 | -97.8976
3.7224 | -97.8683
3.7261 | -97.8601
3.7297 | -97.8777
3.7333 | -97.9262
3.737 | -97.9979
3.7406 | -98.0794
3.7442 | -98.1684
3.7478 | -98.2508
3.7515 | -98.3235
3.7551 | -98.3836
3.7587 | -98.4213
3.7624 | -98.4472
3.766 | -98.4593
3.7696 | -98.4696
3.7732 | -98.4846
3.7769 | -98.4925
3.7805 | -98.5024
3.7841 | -98.5158
3.7878 | -98.5348
3.7914 | -98.5645
3.795 | -98.6041
3.7986 | -98.6589
3.8023 | -98.7251
3.8059 | -98.7971
3.8095 | -98.864
3.8132 | -98.9225
3.8168 | -98.9727
3.8204 | -99.007
3.824 | -99.0156
3.8277 | -99.0123
3.8313 | -99.004
3.8349 | -99.004
3.8385 | -99.0196
3.8422 | -99.0492
3.8458 | -99.0935
3.8494 | -99.1428
3.8531 | -99.1791
3.8567 | -99.1923
3.8603 | -99.1669
3.8639 | -99.1176
3.8676 | -99.0462
3.8712 | -98.9616
3.8748 | -98.8735
3.8785 | -98.7855
3.8821 | -98.7066
3.8857 | -98.6442
3.8893 | -98.5906
3.893 | -98.5472
3.8966 | -98.5158
3.9002 | -98.5127
3.9039 | -98.5226
3.9075 | -98.5492
3.9111 | -98.5965
3.9147 | -98.6635
3.9184 | -98.7497
3.922 | -98.8558
3.9256 | -98.9574
3.9293 | -99.0501
3.9329 | -99.1212
3.9365 | -99.1698
3.9401 | -99.2009
3.9438 | -99.2172
3.9474 | -99.229
3.951 | -99.2551
3.9546 | -99.3025
3.9583 | -99.3677
3.9619 | -99.441
3.9655 | -99.5386
3.9692 | -99.6327
3.9728 | -99.7172
3.9764 | -99.7737
3.98 | -99.7879
3.9837 | -99.7401
3.9873 | -99.64
3.9909 | -99.5125
3.9946 | -99.3888
3.9982 | -99.2886
4.0018 | -99.2379
4.0054 | -99.2458
4.0091 | -99.3043
4.0127 | -99.3914
4.0163 | -99.4808
4.02 | -99.5462
4.0236 | -99.5664
4.0272 | -99.5378
4.0308 | -99.4547
4.0345 | -99.3379
4.0381 | -99.2114
4.0417 | -99.0943
4.0454 | -98.9999
4.049 | -98.9222
4.0526 | -98.8637
4.0562 | -98.8316
4.0599 | -98.8204
4.0635 | -98.8289
4.0671 | -98.856
4.0707 | -98.8918
4.0744 | -98.9208
4.078 | -98.9523
4.0816 | -98.987
4.0853 | -99.0195
4.0889 | -99.0543
4.0925 | -99.0888
4.0961 | -99.1192
4.0998 | -99.1406
4.1034 | -99.1487
4.107 | -99.1592
4.1107 | -99.1795
4.1143 | -99.2067
4.1179 | -99.2504
4.1215 | -99.3133
4.1252 | -99.3836
4.1288 | -99.4607
4.1324 | -99.5293
4.1361 | -99.5659
4.1397 | -99.5516
4.1433 | -99.4962
4.1469 | -99.4152
4.1506 | -99.3125
4.1542 | -99.2186
4.1578 | -99.1487
4.1615 | -99.1115
4.1651 | -99.1068
4.1687 | -99.1246
4.1723 | -99.168
4.176 | -99.2321
4.1796 | -99.3095
4.1832 | -99.4086
4.1868 | -99.4991
4.1905 | -99.5886
4.1941 | -99.6557
4.1977 | -99.7047
4.2014 | -99.7453
4.205 | -99.7666
4.2086 | -99.7771
4.2122 | -99.7865
4.2159 | -99.7977
4.2195 | -99.8225
4.2231 | -99.8493
4.2268 | -99.8851
4.2304 | -99.9373
4.234 | -100.0057
4.2376 | -100.0675
4.2413 | -100.1119
4.2449 | -100.1331
4.2485 | -100.1183
4.2522 | -100.0783
4.2558 | -99.996
4.2594 | -99.8933
4.263 | -99.8
4.2667 | -99.7091
4.2703 | -99.6216
4.2739 | -99.5476
4.2776 | -99.4809
4.2812 | -99.425
4.2848 | -99.3751
4.2884 | -99.3257
4.2921 | -99.2777
4.2957 | -99.2337
4.2993 | -99.1911
4.3029 | -99.1552
4.3066 | -99.1387
4.3102 | -99.1424
4.3138 | -99.159
4.3175 | -99.194
4.3211 | -99.2479
4.3247 | -99.3197
4.3283 | -99.3936
4.332 | -99.4523
4.3356 | -99.5129
4.3392 | -99.5618
4.3429 | -99.5932
4.3465 | -99.6034
4.3501 | -99.5842
4.3537 | -99.5546
4.3574 | -99.5191
4.361 | -99.5016
4.3646 | -99.5082
4.3683 | -99.55
4.3719 | -99.6313
4.3755 | -99.7483
4.3791 | -99.8855
4.3828 | -100.0222
4.3864 | -100.1267
4.39 | -100.1876
4.3937 | -100.2082
4.3973 | -100.1793
4.4009 | -100.1096
4.4045 | -100.0118
4.4082 | -99.9042
4.4118 | -99.7942
4.4154 | -99.6968
4.419 | -99.6156
4.4227 | -99.5523
4.4263 | -99.5107
4.4299 | -99.487
4.4336 | -99.4865
4.4372 | -99.4943
4.4408 | -99.5103
4.4444 | -99.5464
4.4481 | -99.6044
4.4517 | -99.682
4.4553 | -99.7755
4.459 | -99.8765
4.4626 | -99.977
4.4662 | -100.0809
4.4698 | -100.1797
4.4735 | -100.2613
4.4771 | -100.3287
4.4807 | -100.3841
4.4844 | -100.4266
4.488 | -100.4527
4.4916 | -100.4703
4.4952 | -100.4832
4.4989 | -100.4872
4.5025 | -100.4756
4.5061 | -100.4625
4.5098 | -100.4466
4.5134 | -100.4154
4.517 | -100.3817
4.5206 | -100.3474
4.5243 | -100.3209
4.5279 | -100.2967
4.5315 | -100.2765
4.5351 | -100.2601
4.5388 | -100.2526
4.5424 | -100.267
4.546 | -100.2914
4.5497 | -100.318
4.5533 | -100.344
4.5569 | -100.3606
4.5605 | -100.3515
4.5642 | -100.3218
4.5678 | -100.287
4.5714 | -100.245
4.5751 | -100.1945
4.5787 | -100.1423
4.5823 | -100.097
4.5859 | -100.0561
4.5896 | -100.0208
4.5932 | -99.9922
4.5968 | -99.9769
4.6005 | -99.9831
4.6041 | -100.0088
4.6077 | -100.0458
4.6113 | -100.0922
4.615 | -100.1429
4.6186 | -100.1928
4.6222 | -100.2464
4.6258 | -100.302
4.6295 | -100.345
4.6331 | -100.3714
4.6367 | -100.3705
4.6404 | -100.3524
4.644 | -100.3216
4.6476 | -100.2824
4.6512 | -100.2464
4.6549 | -100.2239
4.6585 | -100.2116
4.6621 | -100.2091
4.6658 | -100.2262
4.6694 | -100.2448
4.673 | -100.2684
4.6766 | -100.3036
4.6803 | -100.3418
4.6839 | -100.3928
4.6875 | -100.4462
4.6912 | -100.5008
4.6948 | -100.5592
4.6984 | -100.6231
4.702 | -100.6761
4.7057 | -100.7123
4.7093 | -100.691
4.7129 | -100.6818
4.7166 | -100.655
4.7202 | -100.6134
4.7238 | -100.5562
4.7274 | -100.4947
4.7311 | -100.4385
4.7347 | -100.3776
4.7383 | -100.3189
4.742 | -100.2718
4.7456 | -100.2394
4.7492 | -100.2193
4.7528 | -100.2015
4.7565 | -100.1821
4.7601 | -100.1598
4.7637 | -100.1341
4.7673 | -100.1016
4.771 | -100.0618
4.7746 | -100.0317
4.7782 | -100.02
4.7819 | -100.0207
4.7855 | -100.0362
4.7891 | -100.0647
4.7927 | -100.1054
4.7964 | -100.161
4.8 | -100.2194
4.8036 | -100.2718
4.8073 | -100.314
4.8109 | -100.3466
4.8145 | -100.3566
4.8181 | -100.355
4.8218 | -100.3402
4.8254 | -100.321
4.829 | -100.3103
4.8327 | -100.3106
4.8363 | -100.3225
4.8399 | -100.3438
4.8435 | -100.3823
4.8472 | -100.4307
4.8508 | -100.4691
4.8544 | -100.4916
4.8581 | -100.5014
4.8617 | -100.506
4.8653 | -100.5046
4.8689 | -100.493
4.8726 | -100.4852
4.8762 | -100.4915
4.8798 | -100.5057
4.8834 | -100.5319
4.8871 | -100.5722
4.8907 | -100.6168
4.8943 | -100.6582
4.898 | -100.6988
4.9016 | -100.7346
4.9052 | -100.7685
4.9088 | -100.7924
4.9125 | -100.8054
4.9161 | -100.8086
4.9197 | -100.793
4.9234 | -100.7673
4.927 | -100.7306
4.9306 | -100.6738
4.9342 | -100.6072
4.9379 | -100.5293
4.9415 | -100.4506
4.9451 | -100.3732
4.9488 | -100.2899
4.9524 | -100.2161
4.956 | -100.1557
4.9596 | -100.1113
4.9633 | -100.0785
4.9669 | -100.0678
4.9705 | -100.0765
4.9741 | -100.1045
4.9778 | -100.1483
4.9814 | -100.198
4.985 | -100.2484
4.9887 | -100.2982
4.9923 | -100.3425
4.9959 | -100.3913
4.9995 | -100.4434
5.0032 | -100.5017
5.0068 | -100.5705
5.0104 | -100.6374
5.0141 | -100.7057
5.0177 | -100.7747
5.0213 | -100.8518
5.0249 | -100.9275
5.0286 | -100.9934
5.0322 | -101.0501
5.0358 | -101.0902
5.0395 | -101.1214
5.0431 | -101.1431
5.0467 | -101.1581
5.0503 | -101.1577
5.054 | -101.1592
5.0576 | -101.1641
5.0612 | -101.1708
5.0649 | -101.1755
5.0685 | -101.1765
5.0721 | -101.1719
5.0757 | -101.1683
5.0794 | -101.168
5.083 | -101.1686
5.0866 | -101.1782
5.0902 | -101.1955
5.0939 | -101.2132
5.0975 | -101.2307
5.1011 | -101.2514
5.1048 | -101.2688
5.1084 | -101.288
5.112 | -101.2968
5.1156 | -101.2979
5.1193 | -101.2919
5.1229 | -101.284
5.1265 | -101.2825
5.1302 | -101.2637
5.1338 | -101.2361
5.1374 | -101.1982
5.141 | -101.158
5.1447 | -101.1197
5.1483 | -101.0822
5.1519 | -101.0465
5.1556 | -101.0304
5.1592 | -101.0319
5.1628 | -101.0383
5.1664 | -101.0534
5.1701 | -101.0838
5.1737 | -101.127
5.1773 | -101.1807
5.181 | -101.2359
5.1846 | -101.2838
5.1882 | -101.3313
5.1918 | -101.3532
5.1955 | -101.3399
5.1991 | -101.2938
5.2027 | -101.2174
5.2063 | -101.1262
5.21 | -101.0257
5.2136 | -100.939
5.2172 | -100.8644
5.2209 | -100.814
5.2245 | -100.7989
5.2281 | -100.8085
5.2317 | -100.8593
5.2354 | -100.8855
5.239 | -100.9182
5.2426 | -100.955
5.2463 | -100.9887
5.2499 | -101.018
5.2535 | -101.0461
5.2571 | -101.0723
5.2608 | -101.0962
5.2644 | -101.1169
5.268 | -101.13
5.2717 | -101.145
5.2753 | -101.1541
5.2789 | -101.1587
5.2825 | -101.1531
5.2862 | -101.1394
5.2898 | -101.1259
5.2934 | -101.1009
5.2971 | -101.0789
5.3007 | -101.0648
5.3043 | -101.0543
5.3079 | -101.0553
5.3116 | -101.061
5.3152 | -101.0733
5.3188 | -101.0966
5.3224 | -101.1293
5.3261 | -101.1698
5.3297 | -101.2085
5.3333 | -101.2345
5.337 | -101.2492
5.3406 | -101.2543
5.3442 | -101.2587
5.3478 | -101.2585
5.3515 | -101.2518
5.3551 | -101.239
5.3587 | -101.2329
5.3624 | -101.2292
5.366 | -101.2321
5.3696 | -101.2273
5.3732 | -101.2118
5.3769 | -101.1912
5.3805 | -101.1558
5.3841 | -101.1103
5.3878 | -101.0648
5.3914 | -101.0213
5.395 | -100.9846
5.3986 | -100.9469
5.4023 | -100.921
5.4059 | -100.9035
5.4095 | -100.9
5.4132 | -100.9109
5.4168 | -100.9376
5.4204 | -100.9767
5.424 | -101.027
5.4277 | -101.0796
5.4313 | -101.1386
5.4349 | -101.1988
5.4385 | -101.2553
5.4422 | -101.2968
5.4458 | -101.3225
5.4494 | -101.332
5.4531 | -101.3273
5.4567 | -101.3037
5.4603 | -101.2722
5.4639 | -101.2422
5.4676 | -101.2141
5.4712 | -101.1891
5.4748 | -101.1746
5.4785 | -101.1704
5.4821 | -101.1766
5.4857 | -101.1859
5.4893 | -101.2049
5.493 | -101.2293
5.4966 | -101.2685
5.5002 | -101.3112
5.5039 | -101.3572
5.5075 | -101.4061
5.5111 | -101.4578
5.5147 | -101.507
5.5184 | -101.5514
5.522 | -101.5899
5.5256 | -101.6213
5.5293 | -101.6403
5.5329 | -101.6533
5.5365 | -101.6585
5.5401 | -101.6565
5.5438 | -101.645
5.5474 | -101.6208
5.551 | -101.5905
5.5546 | -101.562
5.5583 | -101.5302
5.5619 | -101.4953
5.5655 | -101.4689
5.5692 | -101.4467
5.5728 | -101.4285
5.5764 | -101.4087
5.58 | -101.3858
5.5837 | -101.3705
5.5873 | -101.3469
5.5909 | -101.3172
5.5946 | -101.2792
5.5982 | -101.2347
5.6018 | -101.1886
5.6054 | -101.1471
5.6091 | -101.1141
5.6127 | -101.0889
5.6163 | -101.0755
5.62 | -101.0806
5.6236 | -101.1004
5.6272 | -101.139
5.6308 | -101.1953
5.6345 | -101.2546
5.6381 | -101.3103
5.6417 | -101.3619
5.6454 | -101.4075
5.649 | -101.4291
5.6526 | -101.4278
5.6562 | -101.4018
5.6599 | -101.3526
5.6635 | -101.28
5.6671 | -101.2018
5.6707 | -101.1247
5.6744 | -101.0473
5.678 | -100.9862
5.6816 | -100.9443
5.6853 | -100.9183
5.6889 | -100.9014
5.6925 | -100.9
5.6961 | -100.9072
5.6998 | -100.9129
5.7034 | -100.9173
5.707 | -100.9158
5.7107 | -100.9044
5.7143 | -100.8845
5.7179 | -100.8633
5.7215 | -100.8458
5.7252 | -100.8302
5.7288 | -100.8128
5.7324 | -100.8037
5.7361 | -100.8021
5.7397 | -100.8086
5.7433 | -100.8235
5.7469 | -100.8399
5.7506 | -100.8507
5.7542 | -100.8393
5.7578 | -100.8298
5.7615 | -100.8205
5.7651 | -100.8037
5.7687 | -100.7773
5.7723 | -100.7523
5.776 | -100.7281
5.7796 | -100.7111
5.7832 | -100.7025
5.7868 | -100.7028
5.7905 | -100.709
5.7941 | -100.7329
5.7977 | -100.7702
5.8014 | -100.8121
5.805 | -100.858
5.8086 | -100.8951
5.8122 | -100.922
5.8159 | -100.9374
5.8195 | -100.9345
5.8231 | -100.9167
5.8268 | -100.8871
5.8304 | -100.8418
5.834 | -100.793
5.8376 | -100.7408
5.8413 | -100.6901
5.8449 | -100.6454
5.8485 | -100.6073
5.8522 | -100.5707
5.8558 | -100.5375
5.8594 | -100.5083
5.863 | -100.482
5.8667 | -100.4569
5.8703 | -100.4313
5.8739 | -100.4148
5.8776 | -100.404
5.8812 | -100.4019
5.8848 | -100.4023
5.8884 | -100.4123
5.8921 | -100.4334
5.8957 | -100.4558
5.8993 | -100.4877
5.9029 | -100.5292
5.9066 | -100.5758
5.9102 | -100.6266
5.9138 | -100.6779
5.9175 | -100.7243
5.9211 | -100.7664
5.9247 | -100.7977
5.9283 | -100.817
5.932 | -100.8302
5.9356 | -100.832
5.9392 | -100.8294
5.9429 | -100.8297
5.9465 | -100.8333
5.9501 | -100.8418
5.9537 | -100.8518
5.9574 | -100.8713
5.961 | -100.894
5.9646 | -100.9168
5.9683 | -100.9433
5.9719 | -100.9668
5.9755 | -100.9926
5.9791 | -101.0231
5.9828 | -101.0577
5.9864 | -101.0959
5.99 | -101.1352
5.9937 | -101.1701
5.9973 | -101.2102
6.0009 | -101.2508
6.0045 | -101.2926
6.0082 | -101.3451
6.0118 | -101.4073
6.0154 | -101.475
6.019 | -101.547
6.0227 | -101.6172
6.0263 | -101.6849
6.0299 | -101.7401
6.0336 | -101.7749
6.0372 | -101.8014
6.0408 | -101.8161
6.0444 | -101.8124
6.0481 | -101.7937
6.0517 | -101.7674
6.0553 | -101.7368
6.059 | -101.7013
6.0626 | -101.6784
6.0662 | -101.6548
6.0698 | -101.6319
6.0735 | -101.6185
6.0771 | -101.6148
6.0807 | -101.6201
6.0844 | -101.6359
6.088 | -101.6453
6.0916 | -101.6604
6.0952 | -101.6746
6.0989 | -101.6759
6.1025 | -101.6697
6.1061 | -101.6588
6.1098 | -101.6424
6.1134 | -101.6225
6.117 | -101.6048
6.1206 | -101.5899
6.1243 | -101.583
6.1279 | -101.5747
6.1315 | -101.5777
6.1351 | -101.5908
6.1388 | -101.6125
6.1424 | -101.644
6.146 | -101.6842
6.1497 | -101.7348
6.1533 | -101.7858
6.1569 | -101.8432
6.1605 | -101.8999
6.1642 | -101.9518
6.1678 | -101.9976
6.1714 | -102.036
6.1751 | -102.0729
6.1787 | -102.113
6.1823 | -102.1489
6.1859 | -102.1816
6.1896 | -102.2125
6.1932 | -102.2412
6.1968 | -102.2652
6.2005 | -102.2811
6.2041 | -102.2833
6.2077 | -102.2745
6.2113 | -102.2587
6.215 | -102.23
6.2186 | -102.1985
6.2222 | -102.1576
6.2259 | -102.1125
6.2295 | -102.0656
6.2331 | -102.0198
6.2367 | -101.9835
6.2404 | -101.9589
6.244 | -101.9464
6.2476 | -101.9492
6.2512 | -101.9619
6.2549 | -101.9885
6.2585 | -102.0246
6.2621 | -102.0683
6.2658 | -102.1114
6.2694 | -102.1535
6.273 | -102.1861
6.2766 | -102.2121
6.2803 | -102.2024
6.2839 | -102.2158
6.2875 | -102.2166
6.2912 | -102.2122
6.2948 | -102.2078
6.2984 | -102.2025
6.302 | -102.1902
6.3057 | -102.1735
6.3093 | -102.1616
6.3129 | -102.1491
6.3166 | -102.1419
6.3202 | -102.1412
6.3238 | -102.1437
6.3274 | -102.1491
6.3311 | -102.1508
6.3347 | -102.1511
6.3383 | -102.1623
6.3419 | -102.1772
6.3456 | -102.2013
6.3492 | -102.2253
6.3528 | -102.2539
6.3565 | -102.2828
6.3601 | -102.3172
6.3637 | -102.3546
6.3673 | -102.3988
6.371 | -102.4457
6.3746 | -102.4893
6.3782 | -102.5245
6.3819 | -102.5547
6.3855 | -102.5828
6.3891 | -102.6046
6.3927 | -102.6167
6.3964 | -102.6272
6.4 | -102.6445
6.4036 | -102.6631
6.4073 | -102.6701
6.4109 | -102.6734
6.4145 | -102.6783
6.4181 | -102.6765
6.4218 | -102.6637
6.4254 | -102.6546
6.429 | -102.6511
6.4327 | -102.649
6.4363 | -102.6388
6.4399 | -102.63
6.4435 | -102.6141
6.4472 | -102.6003
6.4508 | -102.5879
6.4544 | -102.5731
6.4581 | -102.5621
6.4617 | -102.5579
6.4653 | -102.5523
6.4689 | -102.5465
6.4726 | -102.5538
6.4762 | -102.5631
6.4798 | -102.5759
6.4834 | -102.5988
6.4871 | -102.6214
6.4907 | -102.6401
6.4943 | -102.6495
6.498 | -102.6571
6.5016 | -102.6591
6.5052 | -102.6572
6.5088 | -102.6428
6.5125 | -102.6214
6.5161 | -102.602
6.5197 | -102.5851
6.5234 | -102.5581
6.527 | -102.5268
6.5306 | -102.4925
6.5342 | -102.4629
6.5379 | -102.4437
6.5415 | -102.4334
6.5451 | -102.4249
6.5488 | -102.4215
6.5524 | -102.4273
6.556 | -102.4348
6.5596 | -102.4488
6.5633 | -102.4684
6.5669 | -102.4922
6.5705 | -102.515
6.5742 | -102.5419
6.5778 | -102.571
6.5814 | -102.5956
6.585 | -102.6286
6.5887 | -102.6685
6.5923 | -102.7095
6.5959 | -102.7447
6.5995 | -102.7831
6.6032 | -102.8227
6.6068 | -102.8584
6.6104 | -102.8846
6.6141 | -102.9009
6.6177 | -102.9088
6.6213 | -102.9055
6.6249 | -102.8929
6.6286 | -102.8742
6.6322 | -102.8404
6.6358 | -102.7956
6.6395 | -102.7491
6.6431 | -102.6979
6.6467 | -102.6468
6.6503 | -102.601
6.654 | -102.5602
6.6576 | -102.5251
6.6612 | -102.5054
6.6649 | -102.4911
6.6685 | -102.4876
6.6721 | -102.4945
6.6757 | -102.5024
6.6794 | -102.5149
6.683 | -102.5336
6.6866 | -102.5558
6.6902 | -102.5785
6.6939 | -102.603
6.6975 | -102.6287
6.7011 | -102.6502
6.7048 | -102.6678
6.7084 | -102.6833
6.712 | -102.6949
6.7156 | -102.7007
6.7193 | -102.7035
6.7229 | -102.7111
6.7265 | -102.7194
6.7302 | -102.7308
6.7338 | -102.7409
6.7374 | -102.7491
6.741 | -102.7636
6.7447 | -102.7799
6.7483 | -102.7942
6.7519 | -102.8068
6.7556 | -102.8197
6.7592 | -102.8314
6.7628 | -102.8486
6.7664 | -102.8637
6.7701 | -102.8818
6.7737 | -102.9075
6.7773 | -102.9389
6.781 | -102.9629
6.7846 | -102.982
6.7882 | -103.0041
6.7918 | -103.018
6.7955 | -103.0218
6.7991 | -103.0185
6.8027 | -103.0028
6.8063 | -102.9948
6.81 | -102.9806
6.8136 | -102.9706
6.8172 | -102.9732
6.8209 | -102.9799
6.8245 | -102.9869
6.8281 | -103.0031
6.8317 | -103.0219
6.8354 | -103.0476
6.839 | -103.073
6.8426 | -103.0871
6.8463 | -103.0929
6.8499 | -103.0923
6.8535 | -103.087
6.8571 | -103.0717
6.8608 | -103.0468
6.8644 | -103.017
6.868 | -102.9831
6.8717 | -102.9467
6.8753 | -102.9037
6.8789 | -102.86
6.8825 | -102.8173
6.8862 | -102.7763
6.8898 | -102.7319
6.8934 | -102.6928
6.8971 | -102.6594
6.9007 | -102.6295
6.9043 | -102.5997
6.9079 | -102.575
6.9116 | -102.5513
6.9152 | -102.5243
6.9188 | -102.4998
6.9224 | -102.4765
6.9261 | -102.461
6.9297 | -102.4515
6.9333 | -102.4464
6.937 | -102.4495
6.9406 | -102.4589
6.9442 | -102.4793
6.9478 | -102.5034
6.9515 | -102.5305
6.9551 | -102.5632
6.9587 | -102.5959
6.9624 | -102.6297
6.966 | -102.6646
6.9696 | -102.6957
6.9732 | -102.7249
6.9769 | -102.7552
6.9805 | -102.7875
6.9841 | -102.815
6.9878 | -102.8424
6.9914 | -102.8655
6.995 | -102.8916
6.9986 | -102.9176
7.0023 | -102.9486
7.0059 | -102.9833
7.0095 | -103.0239
7.0132 | -103.0677
7.0168 | -103.1156
7.0204 | -103.1688
7.024 | -103.2152
7.0277 | -103.2519
7.0313 | -103.2825
7.0349 | -103.3011
7.0385 | -103.3143
7.0422 | -103.3242
7.0458 | -103.3243
7.0494 | -103.3109
7.0531 | -103.2944
7.0567 | -103.2707
7.0603 | -103.237
7.0639 | -103.2008
7.0676 | -103.1663
7.0712 | -103.1325
7.0748 | -103.1028
7.0785 | -103.0822
7.0821 | -103.066
7.0857 | -103.0638
7.0893 | -103.077
7.093 | -103.1028
7.0966 | -103.1395
7.1002 | -103.1778
7.1039 | -103.2189
7.1075 | -103.2577
7.1111 | -103.2922
7.1147 | -103.3194
7.1184 | -103.3366
7.122 | -103.3406
7.1256 | -103.3353
7.1293 | -103.3225
7.1329 | -103.3108
7.1365 | -103.2959
7.1401 | -103.2804
7.1438 | -103.2542
7.1474 | -103.2245
7.151 | -103.199
7.1546 | -103.1777
7.1583 | -103.1575
7.1619 | -103.1372
7.1655 | -103.115
7.1692 | -103.0948
7.1728 | -103.0734
7.1764 | -103.0501
7.18 | -103.0292
7.1837 | -103.0088
7.1873 | -102.995
7.1909 | -102.9838
7.1946 | -102.9749
7.1982 | -102.9666
7.2018 | -102.9693
7.2054 | -102.9755
7.2091 | -102.9872
7.2127 | -103.0081
7.2163 | -103.0295
7.22 | -103.0508
7.2236 | -103.0784
7.2272 | -103.1049
7.2308 | -103.1241
7.2345 | -103.1373
7.2381 | -103.1323
7.2417 | -103.1207
7.2454 | -103.1101
7.249 | -103.1017
7.2526 | -103.0881
7.2562 | -103.0726
7.2599 | -103.0635
7.2635 | -103.0618
7.2671 | -103.0663
7.2707 | -103.083
7.2744 | -103.1071
7.278 | -103.1423
7.2816 | -103.189
7.2853 | -103.2417
7.2889 | -103.2993
7.2925 | -103.3589
7.2961 | -103.415
7.2998 | -103.464
7.3034 | -103.5141
7.307 | -103.5538
7.3107 | -103.5844
7.3143 | -103.5962
7.3179 | -103.5961
7.3215 | -103.59
7.3252 | -103.5521
7.3288 | -103.5325
7.3324 | -103.5009
7.3361 | -103.4679
7.3397 | -103.4332
7.3433 | -103.3998
7.3469 | -103.3761
7.3506 | -103.3466
7.3542 | -103.3143
7.3578 | -103.2915
7.3615 | -103.2782
7.3651 | -103.2702
7.3687 | -103.2646
7.3723 | -103.263
7.376 | -103.2682
7.3796 | -103.2722
7.3832 | -103.2813
7.3868 | -103.2913
7.3905 | -103.2994
7.3941 | -103.3026
7.3977 | -103.3054
7.4014 | -103.3112
7.405 | -103.3169
7.4086 | -103.3203
7.4122 | -103.3169
7.4159 | -103.3105
7.4195 | -103.3002
7.4231 | -103.2871
7.4268 | -103.2683
7.4304 | -103.246
7.434 | -103.2194
7.4376 | -103.1889
7.4413 | -103.16
7.4449 | -103.13
7.4485 | -103.1039
7.4522 | -103.0821
7.4558 | -103.0641
7.4594 | -103.053
7.463 | -103.0531
7.4667 | -103.063
7.4703 | -103.0747
7.4739 | -103.0918
7.4776 | -103.1157
7.4812 | -103.1449
7.4848 | -103.1788
7.4884 | -103.2177
7.4921 | -103.2657
7.4957 | -103.3127
7.4993 | -103.3615
7.5029 | -103.4131
7.5066 | -103.4703
7.5102 | -103.5259
7.5138 | -103.5733
7.5175 | -103.6198
7.5211 | -103.6602
7.5247 | -103.6948
7.5283 | -103.7161
7.532 | -103.7328
7.5356 | -103.7382
7.5392 | -103.7261
7.5429 | -103.7053
7.5465 | -103.683
7.5501 | -103.6588
7.5537 | -103.6275
7.5574 | -103.594
7.561 | -103.5573
7.5646 | -103.5192
7.5683 | -103.4817
7.5719 | -103.4447
7.5755 | -103.4085
7.5791 | -103.3709
7.5828 | -103.3421
7.5864 | -103.3208
7.59 | -103.305
7.5937 | -103.2953
7.5973 | -103.2897
7.6009 | -103.2876
7.6045 | -103.2924
7.6082 | -103.305
7.6118 | -103.3224
7.6154 | -103.3378
7.619 | -103.3504
7.6227 | -103.3613
7.6263 | -103.3794
7.6299 | -103.3914
7.6336 | -103.4041
7.6372 | -103.4158
7.6408 | -103.4295
7.6444 | -103.4425
7.6481 | -103.4588
7.6517 | -103.4825
7.6553 | -103.5013
7.659 | -103.5187
7.6626 | -103.5354
7.6662 | -103.557
7.6698 | -103.5813
7.6735 | -103.6024
7.6771 | -103.6263
7.6807 | -103.6584
7.6844 | -103.6939
7.688 | -103.7393
7.6916 | -103.7922
7.6952 | -103.8483
7.6989 | -103.9073
7.7025 | -103.9701
7.7061 | -104.0365
7.7098 | -104.107
7.7134 | -104.1744
7.717 | -104.2357
7.7206 | -104.295
7.7243 | -104.3572
7.7279 | -104.4202
7.7315 | -104.4718
7.7351 | -104.5151
7.7388 | -104.5438
7.7424 | -104.5719
7.746 | -104.5876
7.7497 | -104.5856
7.7533 | -104.5757
7.7569 | -104.5557
7.7605 | -104.5324
7.7642 | -104.4989
7.7678 | -104.4625
7.7714 | -104.4149
7.7751 | -104.3713
7.7787 | -104.3251
7.7823 | -104.2745
7.7859 | -104.2303
7.7896 | -104.1891
7.7932 | -104.141
7.7968 | -104.101
7.8005 | -104.0564
7.8041 | -104.0152
7.8077 | -103.9764
7.8113 | -103.9388
7.815 | -103.9034
7.8186 | -103.8715
7.8222 | -103.8372
7.8258 | -103.8089
7.8295 | -103.7868
7.8331 | -103.7628
7.8367 | -103.7425
7.8404 | -103.7314
7.844 | -103.7288
7.8476 | -103.733
7.8512 | -103.7453
7.8549 | -103.7517
7.8585 | -103.7603
7.8621 | -103.7753
7.8658 | -103.7906
7.8694 | -103.8052
7.873 | -103.8181
7.8766 | -103.8254
7.8803 | -103.8331
7.8839 | -103.8428
7.8875 | -103.8476
7.8912 | -103.8437
7.8948 | -103.8408
7.8984 | -103.8401
7.902 | -103.8348
7.9057 | -103.8218
7.9093 | -103.805
7.9129 | -103.7822
7.9166 | -103.761
7.9202 | -103.7339
7.9238 | -103.706
7.9274 | -103.6826
7.9311 | -103.6619
7.9347 | -103.6425
7.9383 | -103.6249
7.942 | -103.6096
7.9456 | -103.5971
7.9492 | -103.5941
7.9528 | -103.5919
7.9565 | -103.594
7.9601 | -103.6019
7.9637 | -103.6143
7.9673 | -103.6294
7.971 | -103.6485
7.9746 | -103.6644
7.9782 | -103.6792
7.9819 | -103.6894
7.9855 | -103.7045
7.9891 | -103.7204
7.9927 | -103.7343
7.9964 | -103.7505
8 | -103.7673
8.0036 | -103.7877
8.0073 | -103.8125
8.0109 | -103.8394
8.0145 | -103.8643
8.0181 | -103.8907
8.0218 | -103.9128
8.0254 | -103.9393
8.029 | -103.9726
8.0327 | -104.0035
8.0363 | -104.033
8.0399 | -104.0591
8.0435 | -104.0828
8.0472 | -104.1013
8.0508 | -104.1113
8.0544 | -104.1149
8.058 | -104.1193
8.0617 | -104.1197
8.0653 | -104.1125
8.0689 | -104.0962
8.0726 | -104.0754
8.0762 | -104.049
8.0798 | -104.0157
8.0834 | -103.9836
8.0871 | -103.9587
8.0907 | -103.9331
8.0943 | -103.9064
8.098 | -103.8855
8.1016 | -103.8747
8.1052 | -103.8675
8.1088 | -103.8586
8.1125 | -103.8569
8.1161 | -103.8568
8.1197 | -103.8545
8.1234 | -103.8565
8.127 | -103.8595
8.1306 | -103.8625
8.1342 | -103.8671
8.1379 | -103.8719
8.1415 | -103.8771
8.1451 | -103.8773
8.1488 | -103.8766
8.1524 | -103.8758
8.156 | -103.8757
8.1596 | -103.8803
8.1633 | -103.8812
8.1669 | -103.8818
8.1705 | -103.8885
8.1741 | -103.8951
8.1778 | -103.9031
8.1814 | -103.9085
8.185 | -103.9132
8.1887 | -103.9216
8.1923 | -103.9316
8.1959 | -103.9428
8.1995 | -103.9535
8.2032 | -103.9594
8.2068 | -103.9632
8.2104 | -103.9713
8.2141 | -103.9828
8.2177 | -103.9914
8.2213 | -103.9946
8.2249 | -103.9978
8.2286 | -104.0046
8.2322 | -104.0076
8.2358 | -104.0054
8.2395 | -104.0073
8.2431 | -104.0026
8.2467 | -104.0017
8.2503 | -104.0002
8.254 | -103.9959
8.2576 | -103.9901
8.2612 | -103.9877
8.2649 | -103.9912
8.2685 | -103.9892
8.2721 | -103.985
8.2757 | -103.9808
8.2794 | -103.9749
8.283 | -103.9706
8.2866 | -103.9672
8.2902 | -103.9591
8.2939 | -103.9548
8.2975 | -103.9476
8.3011 | -103.9404
8.3048 | -103.9421
8.3084 | -103.9416
8.312 | -103.9358
8.3156 | -103.9342
8.3193 | -103.9411
8.3229 | -103.9477
8.3265 | -103.9509
8.3302 | -103.9545
8.3338 | -103.9547
8.3374 | -103.9536
8.341 | -103.9548
8.3447 | -103.9504
8.3483 | -103.9467
8.3519 | -103.9467
8.3556 | -103.9458
8.3592 | -103.9425
8.3628 | -103.9434
8.3664 | -103.946
8.3701 | -103.9511
8.3737 | -103.9588
8.3773 | -103.9676
8.381 | -103.9764
8.3846 | -103.9845
8.3882 | -103.9873
8.3918 | -103.9904
8.3955 | -103.9976
8.3991 | -104.0046
8.4027 | -104.0127
8.4063 | -104.0247
8.41 | -104.0356
8.4136 | -104.0499
8.4172 | -104.0658
8.4209 | -104.0808
8.4245 | -104.0989
8.4281 | -104.1105
8.4317 | -104.1255
8.4354 | -104.141
8.439 | -104.1493
8.4426 | -104.1549
8.4463 | -104.1581
8.4499 | -104.1641
8.4535 | -104.1663
8.4571 | -104.1709
8.4608 | -104.174
8.4644 | -104.171
8.468 | -104.1662
8.4717 | -104.1579
8.4753 | -104.144
8.4789 | -104.1335
8.4825 | -104.1168
8.4862 | -104.0917
8.4898 | -104.062
8.4934 | -104.0252
8.4971 | -103.9846
8.5007 | -103.9445
8.5043 | -103.8977
8.5079 | -103.8502
8.5116 | -103.8014
8.5152 | -103.7553
8.5188 | -103.709
8.5224 | -103.6652
8.5261 | -103.626
8.5297 | -103.5897
8.5333 | -103.5558
8.537 | -103.5344
8.5406 | -103.5206
8.5442 | -103.5135
8.5478 | -103.5124
8.5515 | -103.5168
8.5551 | -103.531
8.5587 | -103.5493
8.5624 | -103.5738
8.566 | -103.6027
8.5696 | -103.6388
8.5732 | -103.6825
8.5769 | -103.7263
8.5805 | -103.7768
8.5841 | -103.828
8.5878 | -103.878
8.5914 | -103.9358
8.595 | -103.997
8.5986 | -104.0556
8.6023 | -104.1096
8.6059 | -104.1531
8.6095 | -104.1958
8.6132 | -104.2358
8.6168 | -104.2681
8.6204 | -104.2983
8.624 | -104.3237
8.6277 | -104.3364
8.6313 | -104.346
8.6349 | -104.3492
8.6385 | -104.3458
8.6422 | -104.3366
8.6458 | -104.3225
8.6494 | -104.3003
8.6531 | -104.2719
8.6567 | -104.2401
8.6603 | -104.2095
8.6639 | -104.1725
8.6676 | -104.1345
8.6712 | -104.0959
8.6748 | -104.0645
8.6785 | -104.0373
8.6821 | -104.0124
8.6857 | -103.9962
8.6893 | -103.9855
8.693 | -103.9803
8.6966 | -103.9765
8.7002 | -103.9768
8.7039 | -103.9816
8.7075 | -103.987
8.7111 | -103.9974
8.7147 | -104.011
8.7184 | -104.0156
8.722 | -104.0115
8.7256 | -104.0077
8.7293 | -104.0014
8.7329 | -103.9926
8.7365 | -103.9781
8.7401 | -103.9601
8.7438 | -103.939
8.7474 | -103.9126
8.751 | -103.8804
8.7546 | -103.8475
8.7583 | -103.8136
8.7619 | -103.784
8.7655 | -103.7559
8.7692 | -103.7277
8.7728 | -103.7056
8.7764 | -103.6869
8.78 | -103.6709
8.7837 | -103.6637
8.7873 | -103.6613
8.7909 | -103.667
8.7946 | -103.6806
8.7982 | -103.7017
8.8018 | -103.7237
8.8054 | -103.746
8.8091 | -103.7693
8.8127 | -103.7945
8.8163 | -103.8205
8.82 | -103.8443
8.8236 | -103.8674
8.8272 | -103.8893
8.8308 | -103.906
8.8345 | -103.9218
8.8381 | -103.9282
8.8417 | -103.9321
8.8454 | -103.9412
8.849 | -103.9452
8.8526 | -103.9426
8.8562 | -103.9436
8.8599 | -103.9439
8.8635 | -103.936
8.8671 | -103.9287
8.8707 | -103.9207
8.8744 | -103.9104
8.878 | -103.9003
8.8816 | -103.888
8.8853 | -103.8742
8.8889 | -103.8593
8.8925 | -103.8451
8.8961 | -103.8368
8.8998 | -103.8278
8.9034 | -103.8236
8.907 | -103.8214
8.9107 | -103.8218
8.9143 | -103.8242
8.9179 | -103.8305
8.9215 | -103.8412
8.9252 | -103.8562
8.9288 | -103.876
8.9324 | -103.8981
8.9361 | -103.9226
8.9397 | -103.9496
8.9433 | -103.9745
8.9469 | -104.0019
8.9506 | -104.0302
8.9542 | -104.0546
8.9578 | -104.0806
8.9615 | -104.1045
8.9651 | -104.1232
8.9687 | -104.1415
8.9723 | -104.157
8.976 | -104.1624
8.9796 | -104.1689
8.9832 | -104.1699
8.9868 | -104.163
8.9905 | -104.1552
8.9941 | -104.1522
8.9977 | -104.1433
9.0014 | -104.133
9.005 | -104.1249
9.0086 | -104.1158
9.0122 | -104.1121
9.0159 | -104.1128
9.0195 | -104.1143
9.0231 | -104.1229
9.0268 | -104.1333
9.0304 | -104.145
9.034 | -104.1633
9.0376 | -104.1801
9.0413 | -104.196
9.0449 | -104.2128
9.0485 | -104.2325
9.0522 | -104.2538
9.0558 | -104.2735
9.0594 | -104.295
9.063 | -104.3118
9.0667 | -104.3322
9.0703 | -104.3538
9.0739 | -104.3758
9.0776 | -104.3977
9.0812 | -104.4151
9.0848 | -104.4306
9.0884 | -104.4411
9.0921 | -104.4515
9.0957 | -104.466
9.0993 | -104.473
9.1029 | -104.4802
9.1066 | -104.4851
9.1102 | -104.4902
9.1138 | -104.489
9.1175 | -104.4869
9.1211 | -104.4886
9.1247 | -104.4885
9.1283 | -104.4818
9.132 | -104.4765
9.1356 | -104.4793
9.1392 | -104.4813
9.1429 | -104.4865
9.1465 | -104.4868
9.1501 | -104.4848
9.1537 | -104.4829
9.1574 | -104.4798
9.161 | -104.4835
9.1646 | -104.4859
9.1683 | -104.4862
9.1719 | -104.4866
9.1755 | -104.4885
9.1791 | -104.4903
9.1828 | -104.4866
9.1864 | -104.4821
9.19 | -104.4754
9.1937 | -104.4668
9.1973 | -104.4636
9.2009 | -104.4576
9.2045 | -104.4486
9.2082 | -104.4376
9.2118 | -104.4291
9.2154 | -104.4194
9.219 | -104.4047
9.2227 | -104.395
9.2263 | -104.3824
9.2299 | -104.37
9.2336 | -104.3617
9.2372 | -104.3489
9.2408 | -104.3289
9.2444 | -104.3059
9.2481 | -104.2859
9.2517 | -104.2637
9.2553 | -104.2362
9.259 | -104.2088
9.2626 | -104.1846
9.2662 | -104.1592
9.2698 | -104.1309
9.2735 | -104.1054
9.2771 | -104.0823
9.2807 | -104.0678
9.2844 | -104.0622
9.288 | -104.0619
9.2916 | -104.0634
9.2952 | -104.0629
9.2989 | -104.0709
9.3025 | -104.083
9.3061 | -104.0956
9.3098 | -104.1092
9.3134 | -104.1246
9.317 | -104.1411
9.3206 | -104.1596
9.3243 | -104.1799
9.3279 | -104.1974
9.3315 | -104.2153
9.3351 | -104.2329
9.3388 | -104.2545
9.3424 | -104.2755
9.346 | -104.2952
9.3497 | -104.3184
9.3533 | -104.3477
9.3569 | -104.3758
9.3605 | -104.4024
9.3642 | -104.4268
9.3678 | -104.4572
9.3714 | -104.4965
9.3751 | -104.5394
9.3787 | -104.5785
9.3823 | -104.6187
9.3859 | -104.6542
9.3896 | -104.684
9.3932 | -104.7113
9.3968 | -104.7373
9.4005 | -104.7538
9.4041 | -104.762
9.4077 | -104.763
9.4113 | -104.7498
9.415 | -104.7261
9.4186 | -104.6804
9.4222 | -104.6486
9.4259 | -104.6136
9.4295 | -104.5752
9.4331 | -104.5386
9.4367 | -104.4961
9.4404 | -104.451
9.444 | -104.405
9.4476 | -104.3592
9.4512 | -104.312
9.4549 | -104.2683
9.4585 | -104.2254
9.4621 | -104.1852
9.4658 | -104.1482
9.4694 | -104.1089
9.473 | -104.0715
9.4766 | -104.0361
9.4803 | -104.0049
9.4839 | -103.98
9.4875 | -103.964
9.4912 | -103.9492
9.4948 | -103.9359
9.4984 | -103.9323
9.502 | -103.9405
9.5057 | -103.9519
9.5093 | -103.9659
9.5129 | -103.9895
9.5166 | -104.0223
9.5202 | -104.0537
9.5238 | -104.0902
9.5274 | -104.1338
9.5311 | -104.1757
9.5347 | -104.215
9.5383 | -104.2579
9.5419 | -104.2955
9.5456 | -104.33
9.5492 | -104.3691
9.5528 | -104.4035
9.5565 | -104.4299
9.5601 | -104.4514
9.5637 | -104.4668
9.5673 | -104.4778
9.571 | -104.4869
9.5746 | -104.4911
9.5782 | -104.4937
9.5819 | -104.4942
9.5855 | -104.4943
9.5891 | -104.4891
9.5927 | -104.481
9.5964 | -104.4694
9.6 | -104.4596
9.6036 | -104.4516
9.6073 | -104.4432
9.6109 | -104.4314
9.6145 | -104.421
9.6181 | -104.4115
9.6218 | -104.4016
9.6254 | -104.3951
9.629 | -104.392
9.6327 | -104.3963
9.6363 | -104.4074
9.6399 | -104.4173
9.6435 | -104.4265
9.6472 | -104.4423
9.6508 | -104.4646
9.6544 | -104.489
9.658 | -104.5154
9.6617 | -104.5474
9.6653 | -104.5821
9.6689 | -104.6162
9.6726 | -104.6516
9.6762 | -104.6847
9.6798 | -104.7124
9.6834 | -104.7414
9.6871 | -104.7697
9.6907 | -104.7907
9.6943 | -104.8099
9.698 | -104.8278
9.7016 | -104.8457
9.7052 | -104.8617
9.7088 | -104.8752
9.7125 | -104.8825
9.7161 | -104.8843
9.7197 | -104.8838
9.7234 | -104.8812
9.727 | -104.8717
9.7306 | -104.8588
9.7342 | -104.8459
9.7379 | -104.8289
9.7415 | -104.8104
9.7451 | -104.7919
9.7488 | -104.7728
9.7524 | -104.7618
9.756 | -104.7477
9.7596 | -104.7309
9.7633 | -104.7166
9.7669 | -104.7
9.7705 | -104.6817
9.7742 | -104.6681
9.7778 | -104.6593
9.7814 | -104.6499
9.785 | -104.6474
9.7887 | -104.6462
9.7923 | -104.6416
9.7959 | -104.6371
9.7995 | -104.6277
9.8032 | -104.6128
9.8068 | -104.5929
9.8104 | -104.5689
9.8141 | -104.5447
9.8177 | -104.5226
9.8213 | -104.4998
9.8249 | -104.4727
9.8286 | -104.4416
9.8322 | -104.411
9.8358 | -104.382
9.8395 | -104.3559
9.8431 | -104.333
9.8467 | -104.3142
9.8503 | -104.3036
9.854 | -104.3005
9.8576 | -104.3026
9.8612 | -104.312
9.8649 | -104.3298
9.8685 | -104.3599
9.8721 | -104.3983
9.8757 | -104.4406
9.8794 | -104.4892
9.883 | -104.5397
9.8866 | -104.5939
9.8902 | -104.6529
9.8939 | -104.7093
9.8975 | -104.7648
9.9011 | -104.8193
9.9048 | -104.8686
9.9084 | -104.9092
9.912 | -104.9417
9.9156 | -104.9622
9.9193 | -104.9746
9.9229 | -104.981
9.9265 | -104.9781
9.9302 | -104.9681
9.9338 | -104.9568
9.9374 | -104.9439
9.941 | -104.9344
9.9447 | -104.9233
9.9483 | -104.9173
9.9519 | -104.9148
9.9556 | -104.9141
9.9592 | -104.9127
9.9628 | -104.9117
9.9664 | -104.9136
9.9701 | -104.9185
9.9737 | -104.9296
9.9773 | -104.9419
9.981 | -104.9522
9.9846 | -104.9658
9.9882 | -104.981
9.9918 | -104.997
9.9955 | -105.0062
9.9991 | -105.0182
10.0027 | -105.028
10.0063 | -105.032
10.01 | -105.0355
10.0136 | -105.0355
10.0172 | -105.0327
10.0209 | -105.0268
10.0245 | -105.0202
10.0281 | -105.0159
10.0317 | -105.0103
10.0354 | -105.0034
10.039 | -104.9968
10.0426 | -104.9945
10.0463 | -104.9879
10.0499 | -104.9833
10.0535 | -104.9815
10.0571 | -104.98
10.0608 | -104.9814
10.0644 | -104.9808
10.068 | -104.9829
10.0717 | -104.982
10.0753 | -104.9844
10.0789 | -104.9858
10.0825 | -104.9867
10.0862 | -104.9884
10.0898 | -104.9891
10.0934 | -104.9895
10.0971 | -104.9845
10.1007 | -104.9782
10.1043 | -104.9759
10.1079 | -104.973
10.1116 | -104.968
10.1152 | -104.9628
10.1188 | -104.9604
10.1224 | -104.9551
10.1261 | -104.9512
10.1297 | -104.9511
10.1333 | -104.9487
10.137 | -104.9442
10.1406 | -104.9392
10.1442 | -104.9421
10.1478 | -104.9433
10.1515 | -104.9458
10.1551 | -104.9464
10.1587 | -104.9473
10.1624 | -104.9498
10.166 | -104.9539
10.1696 | -104.9589
10.1732 | -104.9669
10.1769 | -104.9713
10.1805 | -104.9766
10.1841 | -104.9821
10.1878 | -104.9903
10.1914 | -104.9977
10.195 | -105.0082
10.1986 | -105.0187
10.2023 | -105.0286
10.2059 | -105.0361
10.2095 | -105.0455
10.2132 | -105.0575
10.2168 | -105.069
10.2204 | -105.0809
10.224 | -105.0939
10.2277 | -105.1028
10.2313 | -105.1141
10.2349 | -105.1264
10.2385 | -105.1395
10.2422 | -105.1493
10.2458 | -105.1584
10.2494 | -105.161
10.2531 | -105.1554
10.2567 | -105.1531
10.2603 | -105.1519
10.2639 | -105.1466
10.2676 | -105.1355
10.2712 | -105.1236
10.2748 | -105.1126
10.2785 | -105.1013
10.2821 | -105.0919
10.2857 | -105.0847
10.2893 | -105.0773
10.293 | -105.0648
10.2966 | -105.0548
10.3002 | -105.0485
10.3039 | -105.0476
10.3075 | -105.0427
10.3111 | -105.0363
10.3147 | -105.0281
10.3184 | -105.022
10.322 | -105.0144
10.3256 | -105.0073
10.3293 | -104.9994
10.3329 | -104.9929
10.3365 | -104.9843
10.3401 | -104.9754
10.3438 | -104.9694
10.3474 | -104.9607
10.351 | -104.9555
10.3546 | -104.9532
10.3583 | -104.947
10.3619 | -104.9389
10.3655 | -104.9355
10.3692 | -104.9315
10.3728 | -104.9267
10.3764 | -104.9259
10.38 | -104.9209
10.3837 | -104.9126
10.3873 | -104.9023
10.3909 | -104.8947
10.3946 | -104.8885
10.3982 | -104.8818
10.4018 | -104.8703
10.4054 | -104.8542
10.4091 | -104.8376
10.4127 | -104.8223
10.4163 | -104.8062
10.42 | -104.788
10.4236 | -104.7704
10.4272 | -104.7556
10.4308 | -104.7405
10.4345 | -104.7264
10.4381 | -104.7147
10.4417 | -104.7083
10.4454 | -104.7031
10.449 | -104.6972
10.4526 | -104.6969
10.4562 | -104.7025
10.4599 | -104.7106
10.4635 | -104.7226
10.4671 | -104.7468
10.4707 | -104.7601
10.4744 | -104.7748
10.478 | -104.7957
10.4816 | -104.817
10.4853 | -104.8318
10.4889 | -104.846
10.4925 | -104.8635
10.4961 | -104.8799
10.4998 | -104.8965
10.5034 | -104.9115
10.507 | -104.9237
10.5107 | -104.9372
10.5143 | -104.9525
10.5179 | -104.9663
10.5215 | -104.9853
10.5252 | -105.0017
10.5288 | -105.0226
10.5324 | -105.0467
10.5361 | -105.0734
10.5397 | -105.102
10.5433 | -105.1256
10.5469 | -105.1507
10.5506 | -105.1776
10.5542 | -105.2077
10.5578 | -105.2335
10.5615 | -105.2554
10.5651 | -105.2772
10.5687 | -105.2995
10.5723 | -105.3189
10.576 | -105.3345
10.5796 | -105.3467
10.5832 | -105.3574
10.5868 | -105.3672
10.5905 | -105.3753
10.5941 | -105.3781
10.5977 | -105.3797
10.6014 | -105.3801
10.605 | -105.3731
10.6086 | -105.3644
10.6122 | -105.3538
10.6159 | -105.3423
10.6195 | -105.3271
10.6231 | -105.3083
10.6268 | -105.2954
10.6304 | -105.2813
10.634 | -105.2613
10.6376 | -105.2432
10.6413 | -105.2271
10.6449 | -105.2103
10.6485 | -105.1925
10.6522 | -105.1769
10.6558 | -105.1619
10.6594 | -105.1451
10.663 | -105.1314
10.6667 | -105.1172
10.6703 | -105.1057
10.6739 | -105.0905
10.6776 | -105.0705
10.6812 | -105.055
10.6848 | -105.0429
10.6884 | -105.0313
10.6921 | -105.0173
10.6957 | -105.0019
10.6993 | -104.9888
10.7029 | -104.9725
10.7066 | -104.9601
10.7102 | -104.9471
10.7138 | -104.9326
10.7175 | -104.9205
10.7211 | -104.9083
10.7247 | -104.8944
10.7283 | -104.8807
10.732 | -104.8682
10.7356 | -104.8568
10.7392 | -104.8475
10.7429 | -104.8412
10.7465 | -104.836
10.7501 | -104.8336
10.7537 | -104.8302
10.7574 | -104.8292
10.761 | -104.8324
10.7646 | -104.8405
10.7683 | -104.8524
10.7719 | -104.8672
10.7755 | -104.8853
10.7791 | -104.9068
10.7828 | -104.928
10.7864 | -104.9492
10.79 | -104.975
10.7937 | -105.0029
10.7973 | -105.0365
10.8009 | -105.0643
10.8045 | -105.0893
10.8082 | -105.1161
10.8118 | -105.139
10.8154 | -105.1585
10.819 | -105.1739
10.8227 | -105.1833
10.8263 | -105.1858
10.8299 | -105.1866
10.8336 | -105.184
10.8372 | -105.1737
10.8408 | -105.154
10.8444 | -105.1298
10.8481 | -105.1077
10.8517 | -105.09
10.8553 | -105.0706
10.859 | -105.0527
10.8626 | -105.0341
10.8662 | -105.02
10.8698 | -105.0104
10.8735 | -105.0035
10.8771 | -105.0017
10.8807 | -105.0034
10.8844 | -105.0132
10.888 | -105.024
10.8916 | -105.0367
10.8952 | -105.0562
10.8989 | -105.0771
10.9025 | -105.104
10.9061 | -105.1311
10.9098 | -105.1641
10.9134 | -105.1934
10.917 | -105.2224
10.9206 | -105.252
10.9243 | -105.2791
10.9279 | -105.3012
10.9315 | -105.3235
10.9351 | -105.3443
10.9388 | -105.3633
10.9424 | -105.379
10.946 | -105.3849
10.9497 | -105.3865
10.9533 | -105.3839
10.9569 | -105.3826
10.9605 | -105.3755
10.9642 | -105.3611
10.9678 | -105.3454
10.9714 | -105.3283
10.9751 | -105.307
10.9787 | -105.2872
10.9823 | -105.2685
10.9859 | -105.2489
10.9896 | -105.23
10.9932 | -105.2118
10.9968 | -105.195
11.0005 | -105.1791
11.0041 | -105.1661
11.0077 | -105.1564
11.0113 | -105.1468
11.015 | -105.1406
11.0186 | -105.1367
11.0222 | -105.1346
11.0259 | -105.1369
11.0295 | -105.1453
11.0331 | -105.1567
11.0367 | -105.1689
11.0404 | -105.186
11.044 | -105.2061
11.0476 | -105.2332
11.0512 | -105.2628
11.0549 | -105.2918
11.0585 | -105.3217
11.0621 | -105.351
11.0658 | -105.3828
11.0694 | -105.4126
11.073 | -105.445
11.0766 | -105.4738
11.0803 | -105.4942
11.0839 | -105.512
11.0875 | -105.5273
11.0912 | -105.5382
11.0948 | -105.5428
11.0984 | -105.5396
11.102 | -105.5341
11.1057 | -105.5306
11.1093 | -105.5185
11.1129 | -105.5022
11.1166 | -105.4847
11.1202 | -105.4673
11.1238 | -105.4492
11.1274 | -105.4306
11.1311 | -105.4143
11.1347 | -105.4005
11.1383 | -105.3846
11.1419 | -105.3687
11.1456 | -105.3531
11.1492 | -105.3403
11.1528 | -105.3296
11.1565 | -105.3221
11.1601 | -105.3163
11.1637 | -105.3128
11.1673 | -105.31
11.171 | -105.3044
11.1746 | -105.3015
11.1782 | -105.3023
11.1819 | -105.3022
11.1855 | -105.3052
11.1891 | -105.3091
11.1927 | -105.3103
11.1964 | -105.3114
11.2 | -105.3146
11.2036 | -105.3123
11.2073 | -105.3054
11.2109 | -105.2971
11.2145 | -105.29
11.2181 | -105.2826
11.2218 | -105.2747
11.2254 | -105.2646
11.229 | -105.2503
11.2327 | -105.2335
11.2363 | -105.2155
11.2399 | -105.1936
11.2435 | -105.1747
11.2472 | -105.1565
11.2508 | -105.1386
11.2544 | -105.1192
11.2581 | -105.1019
11.2617 | -105.0858
11.2653 | -105.0661
11.2689 | -105.0452
11.2726 | -105.0275
11.2762 | -105.0127
11.2798 | -104.9959
11.2834 | -104.9798
11.2871 | -104.9682
11.2907 | -104.9585
11.2943 | -104.9486
11.298 | -104.9422
11.3016 | -104.9402
11.3052 | -104.9397
11.3088 | -104.9435
11.3125 | -104.9507
11.3161 | -104.961
11.3197 | -104.9705
11.3234 | -104.9806
11.327 | -104.9913
11.3306 | -105.005
11.3342 | -105.0209
11.3379 | -105.0354
11.3415 | -105.0485
11.3451 | -105.0638
11.3488 | -105.0796
11.3524 | -105.0929
11.356 | -105.1042
11.3596 | -105.1143
11.3633 | -105.1237
11.3669 | -105.1346
11.3705 | -105.1459
11.3742 | -105.1555
11.3778 | -105.1624
11.3814 | -105.1671
11.385 | -105.168
11.3887 | -105.1708
11.3923 | -105.1743
11.3959 | -105.1755
11.3995 | -105.174
11.4032 | -105.1724
11.4068 | -105.1706
11.4104 | -105.1644
11.4141 | -105.158
11.4177 | -105.1527
11.4213 | -105.1436
11.4249 | -105.1363
11.4286 | -105.134
11.4322 | -105.1284
11.4358 | -105.123
11.4395 | -105.114
11.4431 | -105.1076
11.4467 | -105.1074
11.4503 | -105.1063
11.454 | -105.1066
11.4576 | -105.1097
11.4612 | -105.11
11.4649 | -105.1119
11.4685 | -105.1163
11.4721 | -105.1198
11.4757 | -105.1202
11.4794 | -105.1239
11.483 | -105.1306
11.4866 | -105.1369
11.4902 | -105.142
11.4939 | -105.1475
11.4975 | -105.1544
11.5011 | -105.163
11.5048 | -105.1734
11.5084 | -105.1793
11.512 | -105.1802
11.5156 | -105.1844
11.5193 | -105.1899
11.5229 | -105.1964
11.5265 | -105.2036
11.5302 | -105.2084
11.5338 | -105.2097
11.5374 | -105.2108
11.541 | -105.2153
11.5447 | -105.2224
11.5483 | -105.2261
11.5519 | -105.2313
11.5556 | -105.2348
11.5592 | -105.24
11.5628 | -105.2445
11.5664 | -105.252
11.5701 | -105.2587
11.5737 | -105.2636
11.5773 | -105.2684
11.581 | -105.269
11.5846 | -105.2708
11.5882 | -105.276
11.5918 | -105.2787
11.5955 | -105.2774
11.5991 | -105.2783
11.6027 | -105.2807
11.6063 | -105.2839
11.61 | -105.2856
11.6136 | -105.2808
11.6172 | -105.2759
11.6209 | -105.2737
11.6245 | -105.2679
11.6281 | -105.2656
11.6317 | -105.2617
11.6354 | -105.2593
11.639 | -105.2568
11.6426 | -105.2568
11.6463 | -105.2526
11.6499 | -105.2448
11.6535 | -105.2433
11.6571 | -105.2432
11.6608 | -105.2398
11.6644 | -105.2347
11.668 | -105.2298
11.6717 | -105.2247
11.6753 | -105.224
11.6789 | -105.2247
11.6825 | -105.2236
11.6862 | -105.217
11.6898 | -105.2102
11.6934 | -105.2049
11.6971 | -105.1991
11.7007 | -105.1947
11.7043 | -105.192
11.7079 | -105.1856
11.7116 | -105.1759
11.7152 | -105.165
11.7188 | -105.1601
11.7224 | -105.1576
11.7261 | -105.1546
11.7297 | -105.1514
11.7333 | -105.1523
11.737 | -105.1492
11.7406 | -105.1469
11.7442 | -105.145
11.7478 | -105.1465
11.7515 | -105.1481
11.7551 | -105.1484
11.7587 | -105.1454
11.7624 | -105.141
11.766 | -105.1429
11.7696 | -105.1438
11.7732 | -105.1409
11.7769 | -105.1403
11.7805 | -105.1387
11.7841 | -105.1366
11.7878 | -105.1348
11.7914 | -105.1316
11.795 | -105.1279
11.7986 | -105.1227
11.8023 | -105.1193
11.8059 | -105.1149
11.8095 | -105.1108
11.8132 | -105.1077
11.8168 | -105.1078
11.8204 | -105.1108
11.824 | -105.1098
11.8277 | -105.1123
11.8313 | -105.1197
11.8349 | -105.1303
11.8385 | -105.1438
11.8422 | -105.1594
11.8458 | -105.1753
11.8494 | -105.1923
11.8531 | -105.2147
11.8567 | -105.2402
11.8603 | -105.2614
11.8639 | -105.2829
11.8676 | -105.3052
11.8712 | -105.3266
11.8748 | -105.3489
11.8785 | -105.3703
11.8821 | -105.3916
11.8857 | -105.4104
11.8893 | -105.4269
11.893 | -105.444
11.8966 | -105.4598
11.9002 | -105.4779
11.9039 | -105.4921
11.9075 | -105.5081
11.9111 | -105.5272
11.9147 | -105.5433
11.9184 | -105.5552
11.922 | -105.5689
11.9256 | -105.5819
11.9293 | -105.5964
11.9329 | -105.6112
11.9365 | -105.6172
11.9401 | -105.6213
11.9438 | -105.6253
11.9474 | -105.627
11.951 | -105.6269
11.9546 | -105.6269
11.9583 | -105.6294
11.9619 | -105.6313
11.9655 | -105.6313
11.9692 | -105.6327
11.9728 | -105.6293
11.9764 | -105.6259
11.98 | -105.6246
11.9837 | -105.6219
11.9873 | -105.618
11.9909 | -105.6177
11.9946 | -105.6173
11.9982 | -105.6144
12.0018 | -105.6102
12.0054 | -105.6085
12.0091 | -105.6021
12.0127 | -105.6004
12.0163 | -105.6024
12.02 | -105.5985
12.0236 | -105.5917
12.0272 | -105.592
12.0308 | -105.5946
12.0345 | -105.5951
12.0381 | -105.5933
12.0417 | -105.5876
12.0454 | -105.5836
12.049 | -105.5756
12.0526 | -105.5666
12.0562 | -105.5574
12.0599 | -105.5454
12.0635 | -105.531
12.0671 | -105.5138
12.0707 | -105.4931
12.0744 | -105.4705
12.078 | -105.4474
12.0816 | -105.4191
12.0853 | -105.3889
12.0889 | -105.36
12.0925 | -105.3321
12.0961 | -105.3036
12.0998 | -105.2717
12.1034 | -105.2384
12.107 | -105.2071
12.1107 | -105.1776
12.1143 | -105.1508
12.1179 | -105.1263
12.1215 | -105.1014
12.1252 | -105.0797
12.1288 | -105.0584
12.1324 | -105.0348
12.1361 | -105.0176
12.1397 | -105.0026
12.1433 | -104.9892
12.1469 | -104.9814
12.1506 | -104.9733
12.1542 | -104.9683
12.1578 | -104.9641
12.1615 | -104.9648
12.1651 | -104.965
12.1687 | -104.966
12.1723 | -104.9729
12.176 | -104.9818
12.1796 | -104.9925
12.1832 | -105.0099
12.1868 | -105.0243
12.1905 | -105.0374
12.1941 | -105.0557
12.1977 | -105.0727
12.2014 | -105.0855
12.205 | -105.099
12.2086 | -105.1119
12.2122 | -105.1238
12.2159 | -105.1392
12.2195 | -105.1557
12.2231 | -105.1656
12.2268 | -105.1777
12.2304 | -105.1917
12.234 | -105.2034
12.2376 | -105.2201
12.2413 | -105.2355
12.2449 | -105.2518
12.2485 | -105.2692
12.2522 | -105.2851
12.2558 | -105.3027
12.2594 | -105.3229
12.263 | -105.3421
12.2667 | -105.3571
12.2703 | -105.3734
12.2739 | -105.3883
12.2776 | -105.4049
12.2812 | -105.4211
12.2848 | -105.4355
12.2884 | -105.4485
12.2921 | -105.4565
12.2957 | -105.4615
12.2993 | -105.4692
12.3029 | -105.4738
12.3066 | -105.4758
12.3102 | -105.4752
12.3138 | -105.4721
12.3175 | -105.4699
12.3211 | -105.4661
12.3247 | -105.456
12.3283 | -105.4457
12.332 | -105.4354
12.3356 | -105.4258
12.3392 | -105.417
12.3429 | -105.4077
12.3465 | -105.3983
12.3501 | -105.3872
12.3537 | -105.3782
12.3574 | -105.3676
12.361 | -105.3526
12.3646 | -105.338
12.3683 | -105.3293
12.3719 | -105.3241
12.3755 | -105.318
12.3791 | -105.3118
12.3828 | -105.3086
12.3864 | -105.3065
12.39 | -105.3052
12.3937 | -105.303
12.3973 | -105.2964
12.4009 | -105.2959
12.4045 | -105.2971
12.4082 | -105.2942
12.4118 | -105.2936
12.4154 | -105.2914
12.419 | -105.2895
12.4227 | -105.2868
12.4263 | -105.2861
12.4299 | -105.2862
12.4336 | -105.2892
12.4372 | -105.2902
12.4408 | -105.291
12.4444 | -105.2955
12.4481 | -105.2988
12.4517 | -105.3022
12.4553 | -105.3063
12.459 | -105.3156
12.4626 | -105.3236
12.4662 | -105.3311
12.4698 | -105.3437
12.4735 | -105.3554
12.4771 | -105.3666
12.4807 | -105.3815
12.4844 | -105.3982
12.488 | -105.415
12.4916 | -105.428
12.4952 | -105.445
12.4989 | -105.462
12.5025 | -105.4843
12.5061 | -105.5061
12.5098 | -105.5245
12.5134 | -105.5433
12.517 | -105.5597
12.5206 | -105.5761
12.5243 | -105.5945
12.5279 | -105.6119
12.5315 | -105.6257
12.5351 | -105.6359
12.5388 | -105.6506
12.5424 | -105.6674
12.546 | -105.6762
12.5497 | -105.6794
12.5533 | -105.6836
12.5569 | -105.6866
12.5605 | -105.6772
12.5642 | -105.6714
12.5678 | -105.6627
12.5714 | -105.6522
12.5751 | -105.6425
12.5787 | -105.6372
12.5823 | -105.6275
12.5859 | -105.6139
12.5896 | -105.5993
12.5932 | -105.579
12.5968 | -105.561
12.6005 | -105.5427
12.6041 | -105.5259
12.6077 | -105.5097
12.6113 | -105.4942
12.615 | -105.4775
12.6186 | -105.4644
12.6222 | -105.4519
12.6258 | -105.4407
12.6295 | -105.4301
12.6331 | -105.422
12.6367 | -105.413
12.6404 | -105.4032
12.644 | -105.396
12.6476 | -105.3917
12.6512 | -105.3897
12.6549 | -105.3925
12.6585 | -105.3945
12.6621 | -105.3962
12.6658 | -105.4002
12.6694 | -105.4063
12.673 | -105.4149
12.6766 | -105.4186
12.6803 | -105.4224
12.6839 | -105.4262
12.6875 | -105.4285
12.6912 | -105.4273
12.6948 | -105.4262
12.6984 | -105.4289
12.702 | -105.433
12.7057 | -105.4332
12.7093 | -105.4267
12.7129 | -105.4228
12.7166 | -105.4208
12.7202 | -105.4126
12.7238 | -105.4012
12.7274 | -105.3935
12.7311 | -105.3862
12.7347 | -105.376
12.7383 | -105.364
12.7419 | -105.3537
12.7456 | -105.3439
12.7492 | -105.3344
12.7528 | -105.328
12.7565 | -105.3239
12.7601 | -105.317
12.7637 | -105.3107
12.7673 | -105.3082
12.771 | -105.3053
12.7746 | -105.3055
12.7782 | -105.3104
12.7819 | -105.3145
12.7855 | -105.3213
12.7891 | -105.3287
12.7927 | -105.3364
12.7964 | -105.3442
12.8 | -105.3521
12.8036 | -105.3612
12.8073 | -105.3723
12.8109 | -105.3832
12.8145 | -105.3922
12.8181 | -105.4026
12.8218 | -105.4134
12.8254 | -105.4241
12.829 | -105.4362
12.8327 | -105.4477
12.8363 | -105.4545
12.8399 | -105.4599
12.8435 | -105.4689
12.8472 | -105.4786
12.8508 | -105.4853
12.8544 | -105.4909
12.8581 | -105.4971
12.8617 | -105.5052
12.8653 | -105.5098
12.8689 | -105.5143
12.8726 | -105.5197
12.8762 | -105.5274
12.8798 | -105.531
12.8834 | -105.5298
12.8871 | -105.5228
12.8907 | -105.5208
12.8943 | -105.5223
12.898 | -105.5165
12.9016 | -105.5106
12.9052 | -105.5055
12.9088 | -105.5002
12.9125 | -105.489
12.9161 | -105.4784
12.9197 | -105.4723
12.9234 | -105.467
12.927 | -105.4613
12.9306 | -105.4574
12.9342 | -105.4547
12.9379 | -105.4481
12.9415 | -105.4434
12.9451 | -105.4419
12.9488 | -105.4426
12.9524 | -105.4425
12.956 | -105.4453
12.9596 | -105.4508
12.9633 | -105.4575
12.9669 | -105.4652
12.9705 | -105.472
12.9741 | -105.4781
12.9778 | -105.4851
12.9814 | -105.4943
12.985 | -105.5071
12.9887 | -105.5205
12.9923 | -105.5295
12.9959 | -105.5376
12.9995 | -105.5507
13.0032 | -105.5616
13.0068 | -105.5714
13.0104 | -105.5752
13.0141 | -105.5775
13.0177 | -105.5809
13.0213 | -105.5868
13.0249 | -105.5911
13.0286 | -105.5891
13.0322 | -105.5858
13.0358 | -105.5829
13.0395 | -105.5806
13.0431 | -105.5749
13.0467 | -105.5705
13.0503 | -105.5653
13.054 | -105.5549
13.0576 | -105.5458
13.0612 | -105.5366
13.0649 | -105.5279
13.0685 | -105.5183
13.0721 | -105.5121
13.0757 | -105.5036
13.0794 | -105.4935
13.083 | -105.4929
13.0866 | -105.491
13.0902 | -105.4873
13.0939 | -105.4813
13.0975 | -105.4782
13.1011 | -105.4797
13.1048 | -105.4775
13.1084 | -105.4752
13.112 | -105.4756
13.1156 | -105.4767
13.1193 | -105.4766
13.1229 | -105.477
13.1265 | -105.4783
13.1302 | -105.478
13.1338 | -105.4781
13.1374 | -105.4773
13.141 | -105.4731
13.1447 | -105.4645
13.1483 | -105.4527
13.1519 | -105.4411
13.1556 | -105.4292
13.1592 | -105.4181
13.1628 | -105.4087
13.1664 | -105.3981
13.1701 | -105.3843
13.1737 | -105.3699
13.1773 | -105.3522
13.181 | -105.3343
13.1846 | -105.3197
13.1882 | -105.3023
13.1918 | -105.2817
13.1955 | -105.2631
13.1991 | -105.2469
13.2027 | -105.2276
13.2063 | -105.2067
13.21 | -105.1852
13.2136 | -105.1645
13.2172 | -105.1462
13.2209 | -105.129
13.2245 | -105.113
13.2281 | -105.0965
13.2317 | -105.0807
13.2354 | -105.0645
13.239 | -105.0517
13.2426 | -105.0437
13.2463 | -105.0298
13.2499 | -105.0207
13.2535 | -105.013
13.2571 | -105.0084
13.2608 | -105.0074
13.2644 | -105.0057
13.268 | -105.0055
13.2717 | -105.0035
13.2753 | -105.002
13.2789 | -104.9955
13.2825 | -104.9909
13.2862 | -104.9848
13.2898 | -104.9712
13.2934 | -104.9545
13.2971 | -104.938
13.3007 | -104.9197
13.3043 | -104.899
13.3079 | -104.8766
13.3116 | -104.8514
13.3152 | -104.8264
13.3188 | -104.7984
13.3224 | -104.7668
13.3261 | -104.7349
13.3297 | -104.7014
13.3333 | -104.6688
13.337 | -104.6403
13.3406 | -104.6109
13.3442 | -104.58
13.3478 | -104.5482
13.3515 | -104.5174
13.3551 | -104.4898
13.3587 | -104.4646
13.3624 | -104.4443
13.366 | -104.424
13.3696 | -104.4102
13.3732 | -104.4013
13.3769 | -104.3917
13.3805 | -104.3847
13.3841 | -104.3822
13.3878 | -104.3821
13.3914 | -104.3878
13.395 | -104.3935
13.3986 | -104.3998
13.4023 | -104.412
13.4059 | -104.4278
13.4095 | -104.4454
13.4132 | -104.4666
13.4168 | -104.4888
13.4204 | -104.5162
13.424 | -104.5439
13.4277 | -104.5728
13.4313 | -104.6022
13.4349 | -104.6313
13.4385 | -104.6638
13.4422 | -104.6935
13.4458 | -104.7259
13.4494 | -104.7579
13.4531 | -104.7878
13.4567 | -104.8159
13.4603 | -104.845
13.4639 | -104.8725
13.4676 | -104.8972
13.4712 | -104.9251
13.4748 | -104.9507
13.4785 | -104.9775
13.4821 | -105.0045
13.4857 | -105.0291
13.4893 | -105.0485
13.493 | -105.0644
13.4966 | -105.082
13.5002 | -105.0989
13.5039 | -105.1168
13.5075 | -105.1322
13.5111 | -105.1489
13.5147 | -105.1651
13.5184 | -105.1803
13.522 | -105.1948
13.5256 | -105.2098
13.5293 | -105.2217
13.5329 | -105.2335
13.5365 | -105.2465
13.5401 | -105.2601
13.5438 | -105.2741
13.5474 | -105.2845
13.551 | -105.293
13.5546 | -105.3022
13.5583 | -105.3157
13.5619 | -105.3278
13.5655 | -105.3399
13.5692 | -105.3546
13.5728 | -105.3673
13.5764 | -105.3791
13.58 | -105.3913
13.5837 | -105.4044
13.5873 | -105.4204
13.5909 | -105.4333
13.5946 | -105.4392
13.5982 | -105.4485
13.6018 | -105.4561
13.6054 | -105.4605
13.6091 | -105.468
13.6127 | -105.4758
13.6163 | -105.4792
13.62 | -105.4838
13.6236 | -105.4875
13.6272 | -105.4907
13.6308 | -105.4896
13.6345 | -105.4897
13.6381 | -105.4849
13.6417 | -105.4765
13.6454 | -105.47
13.649 | -105.463
13.6526 | -105.4548
13.6562 | -105.4452
13.6599 | -105.4371
13.6635 | -105.4304
13.6671 | -105.4233
13.6707 | -105.4143
13.6744 | -105.4026
13.678 | -105.3932
13.6816 | -105.3874
13.6853 | -105.3856
13.6889 | -105.383
13.6925 | -105.3749
13.6961 | -105.3659
13.6998 | -105.3598
13.7034 | -105.3575
13.707 | -105.3571
13.7107 | -105.3592
13.7143 | -105.3624
13.7179 | -105.3681
13.7215 | -105.3745
13.7252 | -105.3829
13.7288 | -105.3906
13.7324 | -105.3971
13.7361 | -105.4035
13.7397 | -105.4089
13.7433 | -105.4185
13.7469 | -105.431
13.7506 | -105.441
13.7542 | -105.4512
13.7578 | -105.4603
13.7615 | -105.4709
13.7651 | -105.4809
13.7687 | -105.494
13.7723 | -105.5073
13.776 | -105.5187
13.7796 | -105.5257
13.7832 | -105.5331
13.7868 | -105.5447
13.7905 | -105.5549
13.7941 | -105.5601
13.7977 | -105.569
13.8014 | -105.5774
13.805 | -105.5848
13.8086 | -105.5894
13.8122 | -105.5965
13.8159 | -105.6028
13.8195 | -105.6091
13.8231 | -105.6128
13.8268 | -105.6143
13.8304 | -105.6201
13.834 | -105.6246
13.8376 | -105.63
13.8413 | -105.6335
13.8449 | -105.6332
13.8485 | -105.6341
13.8522 | -105.6371
13.8558 | -105.6393
13.8594 | -105.6402
13.863 | -105.6422
13.8667 | -105.6453
13.8703 | -105.6484
13.8739 | -105.6524
13.8776 | -105.6548
13.8812 | -105.6559
13.8848 | -105.6568
13.8884 | -105.659
13.8921 | -105.6654
13.8957 | -105.6769
13.8993 | -105.6872
13.9029 | -105.692
13.9066 | -105.6973
13.9102 | -105.7028
13.9138 | -105.7077
13.9175 | -105.7098
13.9211 | -105.7145
13.9247 | -105.7194
13.9283 | -105.7245
13.932 | -105.73
13.9356 | -105.7384
13.9392 | -105.7466
13.9429 | -105.7508
13.9465 | -105.7512
13.9501 | -105.7519
13.9537 | -105.7489
13.9574 | -105.7475
13.961 | -105.7401
13.9646 | -105.7313
13.9683 | -105.7192
13.9719 | -105.7046
13.9755 | -105.693
13.9791 | -105.6791
13.9828 | -105.6633
13.9864 | -105.6452
13.99 | -105.6264
13.9937 | -105.6047
13.9973 | -105.5816
14.0009 | -105.557
14.0045 | -105.5317
14.0082 | -105.5068
14.0118 | -105.4827
14.0154 | -105.458
14.019 | -105.4322
14.0227 | -105.41
14.0263 | -105.3884
14.0299 | -105.3673
14.0336 | -105.347
14.0372 | -105.329
14.0408 | -105.3157
14.0444 | -105.3078
14.0481 | -105.303
14.0517 | -105.3002
14.0553 | -105.3004
14.059 | -105.3061
14.0626 | -105.3139
14.0662 | -105.3252
14.0698 | -105.3408
14.0735 | -105.3622
14.0771 | -105.3867
14.0807 | -105.4148
14.0844 | -105.4452
14.088 | -105.4771
14.0916 | -105.5103
14.0952 | -105.5406
14.0989 | -105.5701
14.1025 | -105.6044
14.1061 | -105.6404
14.1098 | -105.6766
14.1134 | -105.7127
14.117 | -105.7469
14.1206 | -105.7756
14.1243 | -105.8007
14.1279 | -105.8217
14.1315 | -105.8225
14.1351 | -105.8404
14.1388 | -105.8542
14.1424 | -105.8653
14.146 | -105.8733
14.1497 | -105.8802
14.1533 | -105.8858
14.1569 | -105.8902
14.1605 | -105.8876
14.1642 | -105.8859
14.1678 | -105.8837
14.1714 | -105.8806
14.1751 | -105.8769
14.1787 | -105.8662
14.1823 | -105.8551
14.1859 | -105.8465
14.1896 | -105.8366
14.1932 | -105.8273
14.1968 | -105.8197
14.2005 | -105.8108
14.2041 | -105.8018
14.2077 | -105.7961
14.2113 | -105.7923
14.215 | -105.7879
14.2186 | -105.7869
14.2222 | -105.7897
14.2258 | -105.7914
14.2295 | -105.7943
14.2331 | -105.7972
14.2367 | -105.8025
14.2404 | -105.8103
14.244 | -105.8175
14.2476 | -105.8229
14.2512 | -105.8294
14.2549 | -105.8347
14.2585 | -105.8411
14.2621 | -105.8459
14.2658 | -105.8518
14.2694 | -105.8552
14.273 | -105.8587
14.2766 | -105.8619
14.2803 | -105.8634
14.2839 | -105.8616
14.2875 | -105.8599
14.2912 | -105.8556
14.2948 | -105.8501
14.2984 | -105.8492
14.302 | -105.8465
14.3057 | -105.8418
14.3093 | -105.8339
14.3129 | -105.8216
14.3166 | -105.8114
14.3202 | -105.8037
14.3238 | -105.7955
14.3274 | -105.7875
14.3311 | -105.782
14.3347 | -105.7743
14.3383 | -105.7669
14.342 | -105.7593
14.3456 | -105.7537
14.3492 | -105.7467
14.3528 | -105.7407
14.3565 | -105.7354
14.3601 | -105.7299
14.3637 | -105.7272
14.3673 | -105.7288
14.371 | -105.7297
14.3746 | -105.7292
14.3782 | -105.7289
14.3819 | -105.7283
14.3855 | -105.7299
14.3891 | -105.7305
14.3927 | -105.7297
14.3964 | -105.7327
14.4 | -105.7363
14.4036 | -105.7386
14.4073 | -105.742
14.4109 | -105.7449
14.4145 | -105.7457
14.4181 | -105.7464
14.4218 | -105.7486
14.4254 | -105.7516
14.429 | -105.7515
14.4327 | -105.755
14.4363 | -105.7577
14.4399 | -105.7603
14.4435 | -105.7635
14.4472 | -105.7637
14.4508 | -105.7648
14.4544 | -105.7681
14.4581 | -105.7722
14.4617 | -105.7792
14.4653 | -105.7853
14.4689 | -105.7926
14.4726 | -105.7981
14.4762 | -105.8057
14.4798 | -105.8143
14.4834 | -105.8218
14.4871 | -105.8312
14.4907 | -105.8419
14.4943 | -105.8511
14.498 | -105.8613
14.5016 | -105.87
14.5052 | -105.8773
14.5088 | -105.8854
14.5125 | -105.895
14.5161 | -105.9059
14.5197 | -105.9193
14.5234 | -105.9294
14.527 | -105.9396
14.5306 | -105.9517
14.5342 | -105.9627
14.5379 | -105.9715
14.5415 | -105.9819
14.5451 | -105.99
14.5488 | -105.9978
14.5524 | -106.0019
14.556 | -106.0108
14.5596 | -106.0195
14.5633 | -106.0254
14.5669 | -106.0262
14.5705 | -106.0261
14.5741 | -106.0214
14.5778 | -106.016
14.5814 | -106.0111
14.585 | -106.0053
14.5887 | -105.9986
14.5923 | -105.9875
14.5959 | -105.9785
14.5995 | -105.968
14.6032 | -105.9534
14.6068 | -105.9407
14.6104 | -105.9293
14.6141 | -105.9157
14.6177 | -105.9006
14.6213 | -105.8859
14.6249 | -105.8727
14.6286 | -105.8595
14.6322 | -105.8465
14.6358 | -105.8338
14.6395 | -105.8196
14.6431 | -105.8088
14.6467 | -105.8001
14.6503 | -105.7937
14.654 | -105.7988
14.6576 | -105.7974
14.6612 | -105.7972
14.6649 | -105.7983
14.6685 | -105.8008
14.6721 | -105.8043
14.6757 | -105.8062
14.6794 | -105.8069
14.683 | -105.8099
14.6866 | -105.8127
14.6902 | -105.8139
14.6939 | -105.8167
14.6975 | -105.8192
14.7011 | -105.823
14.7048 | -105.8291
14.7084 | -105.834
14.712 | -105.836
14.7156 | -105.8353
14.7193 | -105.8356
14.7229 | -105.8335
14.7265 | -105.8323
14.7302 | -105.8291
14.7338 | -105.8256
14.7374 | -105.822
14.741 | -105.8172
14.7447 | -105.8135
14.7483 | -105.8092
14.7519 | -105.8027
14.7556 | -105.7944
14.7592 | -105.7861
14.7628 | -105.7808
14.7664 | -105.7707
14.7701 | -105.7618
14.7737 | -105.7557
14.7773 | -105.7529
14.781 | -105.7481
14.7846 | -105.7456
14.7882 | -105.7428
14.7918 | -105.7369
14.7955 | -105.7326
14.7991 | -105.7305
14.8027 | -105.7327
14.8063 | -105.7307
14.81 | -105.7272
14.8136 | -105.7283
14.8172 | -105.7283
14.8209 | -105.7308
14.8245 | -105.7382
14.8281 | -105.7457
14.8317 | -105.7517
14.8354 | -105.757
14.839 | -105.7594
14.8426 | -105.7607
14.8463 | -105.7646
14.8499 | -105.7708
14.8535 | -105.7762
14.8571 | -105.7786
14.8608 | -105.779
14.8644 | -105.7789
14.868 | -105.7802
14.8717 | -105.7796
14.8753 | -105.776
14.8789 | -105.7726
14.8825 | -105.7685
14.8862 | -105.7644
14.8898 | -105.7616
14.8934 | -105.7602
14.8971 | -105.7583
14.9007 | -105.7558
14.9043 | -105.7553
14.9079 | -105.7527
14.9116 | -105.7466
14.9152 | -105.7416
14.9188 | -105.7363
14.9224 | -105.7346
14.9261 | -105.7355
14.9297 | -105.7357
14.9333 | -105.7336
14.937 | -105.7338
14.9406 | -105.7347
14.9442 | -105.736
14.9478 | -105.7368
14.9515 | -105.7363
14.9551 | -105.7346
14.9587 | -105.7319
14.9624 | -105.7295
14.966 | -105.7286
14.9696 | -105.7295
14.9732 | -105.7262
14.9769 | -105.7197
14.9805 | -105.7127
14.9841 | -105.705
14.9878 | -105.6965
14.9914 | -105.6862
14.995 | -105.6746
14.9986 | -105.6617
15.0023 | -105.6494
15.0059 | -105.6397
15.0095 | -105.627
15.0132 | -105.6117
15.0168 | -105.595
15.0204 | -105.5809
15.024 | -105.5644
15.0277 | -105.5424
15.0313 | -105.5226
15.0349 | -105.5015
15.0385 | -105.4787
15.0422 | -105.4605
15.0458 | -105.4422
15.0494 | -105.4235
15.0531 | -105.4083
15.0567 | -105.3958
15.0603 | -105.3826
15.0639 | -105.3669
15.0676 | -105.354
15.0712 | -105.3476
15.0748 | -105.345
15.0785 | -105.3441
15.0821 | -105.3446
15.0857 | -105.3491
15.0893 | -105.3576
15.093 | -105.3662
15.0966 | -105.378
15.1002 | -105.3893
15.1039 | -105.4048
15.1075 | -105.4206
15.1111 | -105.4373
15.1147 | -105.4564
15.1184 | -105.4759
15.122 | -105.4959
15.1256 | -105.5189
15.1293 | -105.5388
15.1329 | -105.5603
15.1365 | -105.5793
15.1401 | -105.5963
15.1438 | -105.6117
15.1474 | -105.625
15.151 | -105.6349
15.1546 | -105.6424
15.1583 | -105.6526
15.1619 | -105.6605
15.1655 | -105.6636
15.1692 | -105.6649
15.1728 | -105.6641
15.1764 | -105.6482
15.18 | -105.6431
15.1837 | -105.6315
15.1873 | -105.6178
15.1909 | -105.6079
15.1946 | -105.5967
15.1982 | -105.582
15.2018 | -105.5643
15.2054 | -105.5466
15.2091 | -105.5276
15.2127 | -105.5079
15.2163 | -105.4894
15.22 | -105.4678
15.2236 | -105.4446
15.2272 | -105.4229
15.2308 | -105.402
15.2345 | -105.3828
15.2381 | -105.3642
15.2417 | -105.3502
15.2454 | -105.338
15.249 | -105.3245
15.2526 | -105.3115
15.2562 | -105.3008
15.2599 | -105.2874
15.2635 | -105.2761
15.2671 | -105.2685
15.2707 | -105.2617
15.2744 | -105.2568
15.278 | -105.2536
15.2816 | -105.2526
15.2853 | -105.254
15.2889 | -105.2514
15.2925 | -105.2493
15.2961 | -105.2503
15.2998 | -105.2546
15.3034 | -105.2575
15.307 | -105.2625
15.3107 | -105.2726
15.3143 | -105.282
15.3179 | -105.2942
15.3215 | -105.307
15.3252 | -105.3199
15.3288 | -105.332
15.3324 | -105.3437
15.3361 | -105.3557
15.3397 | -105.3677
15.3433 | -105.3817
15.3469 | -105.3963
15.3506 | -105.4088
15.3542 | -105.4209
15.3578 | -105.431
15.3615 | -105.4416
15.3651 | -105.4546
15.3687 | -105.4622
15.3723 | -105.467
15.376 | -105.4686
15.3796 | -105.4711
15.3832 | -105.4733
15.3868 | -105.4733
15.3905 | -105.4749
15.3941 | -105.474
15.3977 | -105.4688
15.4014 | -105.4637
15.405 | -105.4623
15.4086 | -105.4575
15.4122 | -105.4472
15.4159 | -105.4381
15.4195 | -105.4329
15.4231 | -105.4277
15.4268 | -105.4191
15.4304 | -105.4094
15.434 | -105.4017
15.4376 | -105.3935
15.4413 | -105.3852
15.4449 | -105.3769
15.4485 | -105.3696
15.4522 | -105.3644
15.4558 | -105.3572
15.4594 | -105.35
15.463 | -105.3399
15.4667 | -105.3326
15.4703 | -105.3284
15.4739 | -105.3272
15.4776 | -105.3264
15.4812 | -105.3298
15.4848 | -105.336
15.4884 | -105.34
15.4921 | -105.3432
15.4957 | -105.3469
15.4993 | -105.3526
15.5029 | -105.3614
15.5066 | -105.3691
15.5102 | -105.3797
15.5138 | -105.3936
15.5175 | -105.4071
15.5211 | -105.4187
15.5247 | -105.432
15.5283 | -105.4444
15.532 | -105.4549
15.5356 | -105.4665
15.5392 | -105.4773
15.5429 | -105.486
15.5465 | -105.4923
15.5501 | -105.4998
15.5537 | -105.5052
15.5574 | -105.5056
15.561 | -105.5097
15.5646 | -105.5107
15.5683 | -105.5119
15.5719 | -105.5122
15.5755 | -105.5098
15.5791 | -105.509
15.5828 | -105.5071
15.5864 | -105.502
15.59 | -105.4982
15.5937 | -105.4955
15.5973 | -105.4916
15.6009 | -105.4838
15.6045 | -105.4751
15.6082 | -105.468
15.6118 | -105.4588
15.6154 | -105.4487
15.619 | -105.4429
15.6227 | -105.4345
15.6263 | -105.4251
15.6299 | -105.4161
15.6336 | -105.4101
15.6372 | -105.4027
15.6408 | -105.3923
15.6444 | -105.3799
15.6481 | -105.3706
15.6517 | -105.3633
15.6553 | -105.3585
15.659 | -105.3564
15.6626 | -105.3547
15.6662 | -105.3531
15.6698 | -105.3516
15.6735 | -105.3489
15.6771 | -105.3467
15.6807 | -105.3478
15.6844 | -105.3521
15.688 | -105.3548
15.6916 | -105.3553
15.6952 | -105.3569
15.6989 | -105.3595
15.7025 | -105.3654
15.7061 | -105.3676
15.7098 | -105.3703
15.7134 | -105.3757
15.717 | -105.3831
15.7206 | -105.3886
15.7243 | -105.3952
15.7279 | -105.4008
15.7315 | -105.4061
15.7351 | -105.4126
15.7388 | -105.4184
15.7424 | -105.4246
15.746 | -105.4349
15.7497 | -105.4463
15.7533 | -105.455
15.7569 | -105.4617
15.7605 | -105.4709
15.7642 | -105.4823
15.7678 | -105.4927
15.7714 | -105.5055
15.7751 | -105.519
15.7787 | -105.5334
15.7823 | -105.5486
15.7859 | -105.5563
15.7896 | -105.5621
15.7932 | -105.5672
15.7968 | -105.5708
15.8005 | -105.5773
15.8041 | -105.5816
15.8077 | -105.5833
15.8113 | -105.5851
15.815 | -105.5873
15.8186 | -105.5839
15.8222 | -105.5784
15.8259 | -105.5749
15.8295 | -105.5719
15.8331 | -105.5628
15.8367 | -105.5487
15.8404 | -105.5363
15.844 | -105.5228
15.8476 | -105.5096
15.8512 | -105.4937
15.8549 | -105.4748
15.8585 | -105.4542
15.8621 | -105.433
15.8658 | -105.4127
15.8694 | -105.3931
15.873 | -105.3719
15.8766 | -105.351
15.8803 | -105.3264
15.8839 | -105.3015
15.8875 | -105.2814
15.8912 | -105.2624
15.8948 | -105.2418
15.8984 | -105.222
15.902 | -105.2046
15.9057 | -105.1876
15.9093 | -105.1739
15.9129 | -105.1587
15.9166 | -105.1482
15.9202 | -105.1398
15.9238 | -105.1339
15.9274 | -105.1299
15.9311 | -105.1271
15.9347 | -105.1281
15.9383 | -105.1294
15.942 | -105.1309
15.9456 | -105.1344
15.9492 | -105.1412
15.9528 | -105.1471
15.9565 | -105.1548
15.9601 | -105.1638
15.9637 | -105.1739
15.9673 | -105.1833
15.971 | -105.1923
15.9746 | -105.2004
15.9782 | -105.2098
15.9819 | -105.2213
15.9855 | -105.2325
15.9891 | -105.2415
15.9927 | -105.2487
15.9964 | -105.2564
16 | -105.2624
16.0036 | -105.2653
16.0073 | -105.266
16.0109 | -105.2688
16.0145 | -105.2728
16.0181 | -105.2771
16.0218 | -105.2781
16.0254 | -105.2774
16.029 | -105.2734
16.0327 | -105.2706
16.0363 | -105.2651
16.0399 | -105.2598
16.0435 | -105.2539
16.0472 | -105.249
16.0508 | -105.2481
16.0544 | -105.2451
16.0581 | -105.2405
16.0617 | -105.2364
16.0653 | -105.2322
16.0689 | -105.2279
16.0726 | -105.2232
16.0762 | -105.2161
16.0798 | -105.2113
16.0834 | -105.2091
16.0871 | -105.2054
16.0907 | -105.2015
16.0943 | -105.1998
16.098 | -105.1973
16.1016 | -105.1963
16.1052 | -105.1963
16.1088 | -105.1949
16.1125 | -105.1932
16.1161 | -105.1911
16.1197 | -105.1901
16.1234 | -105.1906
16.127 | -105.1909
16.1306 | -105.193
16.1342 | -105.1921
16.1379 | -105.1915
16.1415 | -105.1945
16.1451 | -105.1926
16.1488 | -105.1902
16.1524 | -105.1886
16.156 | -105.1892
16.1596 | -105.1907
16.1633 | -105.1906
16.1669 | -105.1896
16.1705 | -105.1875
16.1742 | -105.1857
16.1778 | -105.1826
16.1814 | -105.1793
16.185 | -105.175
16.1887 | -105.1715
16.1923 | -105.1654
16.1959 | -105.1601
16.1995 | -105.1566
16.2032 | -105.1509
16.2068 | -105.146
16.2104 | -105.1411
16.2141 | -105.1357
16.2177 | -105.1318
16.2213 | -105.1306
16.2249 | -105.1367
16.2286 | -105.1363
16.2322 | -105.1389
16.2358 | -105.144
16.2395 | -105.151
16.2431 | -105.1586
16.2467 | -105.1656
16.2503 | -105.173
16.254 | -105.1849
16.2576 | -105.1985
16.2612 | -105.2147
16.2649 | -105.2311
16.2685 | -105.25
16.2721 | -105.2665
16.2757 | -105.2825
16.2794 | -105.3011
16.283 | -105.3194
16.2866 | -105.34
16.2902 | -105.359
16.2939 | -105.3802
16.2975 | -105.4022
16.3011 | -105.4218
16.3048 | -105.4409
16.3084 | -105.4608
16.312 | -105.4784
16.3156 | -105.494
16.3193 | -105.5091
16.3229 | -105.5193
16.3265 | -105.5258
16.3302 | -105.5355
16.3338 | -105.5401
16.3374 | -105.5416
16.341 | -105.5458
16.3447 | -105.5485
16.3483 | -105.5454
16.3519 | -105.5408
16.3556 | -105.5365
16.3592 | -105.5319
16.3628 | -105.5229
16.3664 | -105.5141
16.3701 | -105.5001
16.3737 | -105.4828
16.3773 | -105.468
16.381 | -105.4515
16.3846 | -105.4345
16.3882 | -105.4152
16.3918 | -105.3966
16.3955 | -105.3771
16.3991 | -105.3598
16.4027 | -105.3441
16.4063 | -105.3288
16.41 | -105.3124
16.4136 | -105.2952
16.4172 | -105.2793
16.4209 | -105.2633
16.4245 | -105.2509
16.4281 | -105.2412
16.4317 | -105.2296
16.4354 | -105.2195
16.439 | -105.2104
16.4426 | -105.2051
16.4463 | -105.2013
16.4499 | -105.1959
16.4535 | -105.1928
16.4571 | -105.1917
16.4608 | -105.1958
16.4644 | -105.2004
16.468 | -105.2027
16.4717 | -105.2071
16.4753 | -105.2105
16.4789 | -105.2147
16.4825 | -105.2223
16.4862 | -105.2298
16.4898 | -105.2354
16.4934 | -105.2422
16.4971 | -105.2478
16.5007 | -105.2553
16.5043 | -105.2664
16.5079 | -105.2775
16.5116 | -105.2887
16.5152 | -105.296
16.5188 | -105.3011
16.5224 | -105.3068
16.5261 | -105.3152
16.5297 | -105.3228
16.5333 | -105.3301
16.537 | -105.3364
16.5406 | -105.3404
16.5442 | -105.3438
16.5478 | -105.3471
16.5515 | -105.3507
16.5551 | -105.3561
16.5587 | -105.3646
16.5624 | -105.3703
16.566 | -105.3715
16.5696 | -105.3731
16.5732 | -105.3746
16.5769 | -105.3776
16.5805 | -105.3813
16.5841 | -105.3854
16.5878 | -105.3872
16.5914 | -105.3885
16.595 | -105.3914
16.5986 | -105.3941
16.6023 | -105.3981
16.6059 | -105.4031
16.6095 | -105.4089
16.6132 | -105.4159
16.6168 | -105.4242
16.6204 | -105.4314
16.624 | -105.4376
16.6277 | -105.4453
16.6313 | -105.4556
16.6349 | -105.4662
16.6385 | -105.477
16.6422 | -105.486
16.6458 | -105.4998
16.6494 | -105.5164
16.6531 | -105.53
16.6567 | -105.5445
16.6603 | -105.5584
16.6639 | -105.5711
16.6676 | -105.5879
16.6712 | -105.6057
16.6748 | -105.6207
16.6785 | -105.6359
16.6821 | -105.6521
16.6857 | -105.6696
16.6893 | -105.6843
16.693 | -105.6975
16.6966 | -105.7111
16.7002 | -105.7272
16.7039 | -105.7448
16.7075 | -105.759
16.7111 | -105.7757
16.7147 | -105.7913
16.7184 | -105.8061
16.722 | -105.8204
16.7256 | -105.8334
16.7293 | -105.8457
16.7329 | -105.858
16.7365 | -105.868
16.7401 | -105.8756
16.7438 | -105.885
16.7474 | -105.8901
16.751 | -105.8979
16.7546 | -105.9085
16.7583 | -105.9158
16.7619 | -105.9216
16.7655 | -105.9239
16.7692 | -105.9252
16.7728 | -105.9267
16.7764 | -105.9278
16.78 | -105.9275
16.7837 | -105.9237
16.7873 | -105.9203
16.7909 | -105.9178
16.7946 | -105.9105
16.7982 | -105.9027
16.8018 | -105.8961
16.8054 | -105.8883
16.8091 | -105.8788
16.8127 | -105.8656
16.8163 | -105.8502
16.82 | -105.8366
16.8236 | -105.8234
16.8272 | -105.8084
16.8308 | -105.794
16.8345 | -105.7795
16.8381 | -105.7651
16.8417 | -105.7497
16.8454 | -105.7339
16.849 | -105.7201
16.8526 | -105.7043
16.8562 | -105.6852
16.8599 | -105.6682
16.8635 | -105.6538
16.8671 | -105.6392
16.8707 | -105.6229
16.8744 | -105.6088
16.878 | -105.5977
16.8816 | -105.5861
16.8853 | -105.5739
16.8889 | -105.5646
16.8925 | -105.5569
16.8961 | -105.5515
16.8998 | -105.5459
16.9034 | -105.5422
16.907 | -105.5381
16.9107 | -105.5366
16.9143 | -105.5392
16.9179 | -105.5395
16.9215 | -105.5422
16.9252 | -105.5472
16.9288 | -105.5531
16.9324 | -105.5612
16.9361 | -105.5659
16.9397 | -105.5709
16.9433 | -105.5785
16.9469 | -105.587
16.9506 | -105.595
16.9542 | -105.6028
16.9578 | -105.6115
16.9615 | -105.6185
16.9651 | -105.6246
16.9687 | -105.6306
16.9723 | -105.6354
16.976 | -105.6406
16.9796 | -105.6454
16.9832 | -105.6487
16.9868 | -105.6533
16.9905 | -105.658
16.9941 | -105.6605
16.9977 | -105.6604
17.0014 | -105.6592
17.005 | -105.6573
17.0086 | -105.6583
17.0122 | -105.6557
17.0159 | -105.6502
17.0195 | -105.6444
17.0231 | -105.6388
17.0268 | -105.6343
17.0304 | -105.6261
17.034 | -105.6158
17.0376 | -105.6086
17.0413 | -105.6002
17.0449 | -105.5912
17.0485 | -105.5821
17.0522 | -105.574
17.0558 | -105.5669
17.0594 | -105.5575
17.063 | -105.5486
17.0667 | -105.5412
17.0703 | -105.5362
17.0739 | -105.5284
17.0776 | -105.5187
17.0812 | -105.5104
17.0848 | -105.5028
17.0884 | -105.4941
17.0921 | -105.4861
17.0957 | -105.4819
17.0993 | -105.4799
17.1029 | -105.4776
17.1066 | -105.4738
17.1102 | -105.4696
17.1138 | -105.4669
17.1175 | -105.4624
17.1211 | -105.4576
17.1247 | -105.4565
17.1283 | -105.4543
17.132 | -105.4496
17.1356 | -105.4468
17.1392 | -105.44
17.1429 | -105.4355
17.1465 | -105.4312
17.1501 | -105.426
17.1537 | -105.4191
17.1574 | -105.4122
17.161 | -105.4068
17.1646 | -105.4
17.1683 | -105.3921
17.1719 | -105.3861
17.1755 | -105.3812
17.1791 | -105.3779
17.1828 | -105.3723
17.1864 | -105.365
17.19 | -105.3594
17.1937 | -105.3561
17.1973 | -105.3494
17.2009 | -105.3434
17.2045 | -105.3389
17.2082 | -105.3362
17.2118 | -105.3329
17.2154 | -105.3308
17.219 | -105.3276
17.2227 | -105.3254
17.2263 | -105.3251
17.2299 | -105.3229
17.2336 | -105.3213
17.2372 | -105.3223
17.2408 | -105.324
17.2444 | -105.3287
17.2481 | -105.3315
17.2517 | -105.3342
17.2553 | -105.3397
17.259 | -105.3448
17.2626 | -105.3498
17.2662 | -105.3557
17.2698 | -105.3592
17.2735 | -105.362
17.2771 | -105.3637
17.2807 | -105.3662
17.2844 | -105.367
17.288 | -105.3691
17.2916 | -105.3722
17.2952 | -105.3754
17.2989 | -105.3803
17.3025 | -105.3841
17.3061 | -105.3871
17.3098 | -105.3873
17.3134 | -105.3879
17.317 | -105.3904
17.3206 | -105.3926
17.3243 | -105.393
17.3279 | -105.3951
17.3315 | -105.397
17.3351 | -105.4032
17.3388 | -105.4073
17.3424 | -105.4092
17.346 | -105.4114
17.3497 | -105.414
17.3533 | -105.4182
17.3569 | -105.4235
17.3605 | -105.4283
17.3642 | -105.4328
17.3678 | -105.4369
17.3714 | -105.4393
17.3751 | -105.4397
17.3787 | -105.4428
17.3823 | -105.4471
17.3859 | -105.453
17.3896 | -105.4556
17.3932 | -105.4583
17.3968 | -105.4603
17.4005 | -105.4627
17.4041 | -105.4647
17.4077 | -105.4651
17.4113 | -105.4646
17.415 | -105.4639
17.4186 | -105.4611
17.4222 | -105.4566
17.4259 | -105.4505
17.4295 | -105.4466
17.4331 | -105.4434
17.4367 | -105.4385
17.4404 | -105.4314
17.444 | -105.4267
17.4476 | -105.4199
17.4512 | -105.4111
17.4549 | -105.4023
17.4585 | -105.3949
17.4621 | -105.3823
17.4658 | -105.3691
17.4694 | -105.3548
17.473 | -105.3413
17.4766 | -105.3314
17.4803 | -105.322
17.4839 | -105.3104
17.4875 | -105.2999
17.4912 | -105.2912
17.4948 | -105.2831
17.4984 | -105.2752
17.502 | -105.267
17.5057 | -105.2592
17.5093 | -105.2515
17.5129 | -105.2435
17.5166 | -105.2364
17.5202 | -105.2305
17.5238 | -105.2252
17.5274 | -105.2238
17.5311 | -105.2209
17.5347 | -105.2174
17.5383 | -105.2158
17.5419 | -105.2133
17.5456 | -105.2105
17.5492 | -105.2071
17.5528 | -105.2055
17.5565 | -105.2044
17.5601 | -105.2025
17.5637 | -105.2006
17.5673 | -105.1967
17.571 | -105.1924
17.5746 | -105.1913
17.5782 | -105.1902
17.5819 | -105.1888
17.5855 | -105.1856
17.5891 | -105.1824
17.5927 | -105.18
17.5964 | -105.1805
17.6 | -105.1807
17.6036 | -105.18
17.6073 | -105.179
17.6109 | -105.1796
17.6145 | -105.1799
17.6181 | -105.1795
17.6218 | -105.1796
17.6254 | -105.1819
17.629 | -105.1812
17.6327 | -105.1757
17.6363 | -105.1718
17.6399 | -105.1704
17.6435 | -105.1693
17.6472 | -105.1698
17.6508 | -105.1697
17.6544 | -105.1684
17.658 | -105.1687
17.6617 | -105.1687
17.6653 | -105.1661
17.6689 | -105.1635
17.6726 | -105.1622
17.6762 | -105.1626
17.6798 | -105.1612
17.6834 | -105.1596
17.6871 | -105.1585
17.6907 | -105.1588
17.6943 | -105.1576
17.698 | -105.1544
17.7016 | -105.1533
17.7052 | -105.1528
17.7088 | -105.1493
17.7125 | -105.1503
17.7161 | -105.15
17.7197 | -105.1512
17.7234 | -105.1501
17.727 | -105.1456
17.7306 | -105.1443
17.7342 | -105.1448
17.7379 | -105.1442
17.7415 | -105.1423
17.7451 | -105.1409
17.7488 | -105.1395
17.7524 | -105.1395
17.756 | -105.1395
17.7596 | -105.1407
17.7633 | -105.1415
17.7669 | -105.1427
17.7705 | -105.1436
17.7741 | -105.1444
17.7778 | -105.1456
17.7814 | -105.1489
17.785 | -105.1523
17.7887 | -105.1539
17.7923 | -105.1554
17.7959 | -105.1615
17.7995 | -105.166
17.8032 | -105.1708
17.8068 | -105.1757
17.8104 | -105.1802
17.8141 | -105.1844
17.8177 | -105.1906
17.8213 | -105.198
17.8249 | -105.2069
17.8286 | -105.2142
17.8322 | -105.2209
17.8358 | -105.2286
17.8395 | -105.235
17.8431 | -105.2419
17.8467 | -105.251
17.8503 | -105.2579
17.854 | -105.2649
17.8576 | -105.2712
17.8612 | -105.2783
17.8649 | -105.2864
17.8685 | -105.2947
17.8721 | -105.3043
17.8757 | -105.3122
17.8794 | -105.3205
17.883 | -105.3277
17.8866 | -105.3371
17.8902 | -105.3456
17.8939 | -105.35
17.8975 | -105.3537
17.9011 | -105.3554
17.9048 | -105.3567
17.9084 | -105.3559
17.912 | -105.3524
17.9156 | -105.3476
17.9193 | -105.3464
17.9229 | -105.3428
17.9265 | -105.337
17.9302 | -105.3304
17.9338 | -105.3224
17.9374 | -105.3118
17.941 | -105.3016
17.9447 | -105.2901
17.9483 | -105.2785
17.9519 | -105.2676
17.9556 | -105.2558
17.9592 | -105.2414
17.9628 | -105.229
17.9664 | -105.2148
17.9701 | -105.2001
17.9737 | -105.1829
17.9773 | -105.1661
17.981 | -105.1514
17.9846 | -105.1355
17.9882 | -105.1215
17.9918 | -105.1106
17.9955 | -105.0986
17.9991 | -105.0886
18.0027 | -105.0801
18.0063 | -105.0727
18.01 | -105.0661
18.0136 | -105.0572
18.0172 | -105.0516
18.0209 | -105.0466
18.0245 | -105.0419
18.0281 | -105.039
18.0317 | -105.0334
18.0354 | -105.0276
18.039 | -105.0232
18.0426 | -105.0192
18.0463 | -105.0176
18.0499 | -105.0153
18.0535 | -105.0118
18.0571 | -105.0102
18.0608 | -105.0094
18.0644 | -105.0078
18.068 | -105.006
18.0717 | -105.0042
18.0753 | -105.0017
18.0789 | -104.996
18.0825 | -104.9913
18.0862 | -104.9884
18.0898 | -104.9844
18.0934 | -104.9803
18.0971 | -104.9776
18.1007 | -104.9721
18.1043 | -104.964
18.1079 | -104.9578
18.1116 | -104.9516
18.1152 | -104.9468
18.1188 | -104.9415
18.1224 | -104.9361
18.1261 | -104.9301
18.1297 | -104.9222
18.1333 | -104.9147
18.137 | -104.9044
18.1406 | -104.8966
18.1442 | -104.8896
18.1478 | -104.8819
18.1515 | -104.8756
18.1551 | -104.8694
18.1587 | -104.8644
18.1624 | -104.86
18.166 | -104.8574
18.1696 | -104.857
18.1732 | -104.8558
18.1769 | -104.8492
18.1805 | -104.845
18.1841 | -104.841
18.1878 | -104.8408
18.1914 | -104.8423
18.195 | -104.8437
18.1986 | -104.8447
18.2023 | -104.8457
18.2059 | -104.8486
18.2095 | -104.8519
18.2132 | -104.8547
18.2168 | -104.8601
18.2204 | -104.8647
18.224 | -104.8685
18.2277 | -104.8714
18.2313 | -104.8737
18.2349 | -104.8772
18.2385 | -104.88
18.2422 | -104.8807
18.2458 | -104.8813
18.2494 | -104.8802
18.2531 | -104.8802
18.2567 | -104.8805
18.2603 | -104.8774
18.2639 | -104.8743
18.2676 | -104.8718
18.2712 | -104.8677
18.2748 | -104.8607
18.2785 | -104.8512
18.2821 | -104.8407
18.2857 | -104.8307
18.2893 | -104.8208
18.293 | -104.8107
18.2966 | -104.8003
18.3002 | -104.7894
18.3039 | -104.7767
18.3075 | -104.7608
18.3111 | -104.7439
18.3147 | -104.7252
18.3184 | -104.7046
18.322 | -104.6873
18.3256 | -104.6718
18.3293 | -104.6556
18.3329 | -104.639
18.3365 | -104.6211
18.3401 | -104.6067
18.3438 | -104.5914
18.3474 | -104.5737
18.351 | -104.5575
18.3546 | -104.5399
18.3583 | -104.5242
18.3619 | -104.5103
18.3655 | -104.4967
18.3692 | -104.4831
18.3728 | -104.4715
18.3764 | -104.4608
18.38 | -104.4493
18.3837 | -104.4396
18.3873 | -104.4296
18.3909 | -104.4206
18.3946 | -104.413
18.3982 | -104.4104
18.4018 | -104.405
18.4054 | -104.398
18.4091 | -104.392
18.4127 | -104.3881
18.4163 | -104.385
18.42 | -104.3821
18.4236 | -104.3789
18.4272 | -104.3774
18.4308 | -104.3733
18.4345 | -104.3716
18.4381 | -104.3706
18.4417 | -104.3676
18.4454 | -104.3653
18.449 | -104.3606
18.4526 | -104.3556
18.4562 | -104.3509
18.4599 | -104.3468
18.4635 | -104.3443
18.4671 | -104.3405
18.4707 | -104.3366
18.4744 | -104.3317
18.478 | -104.326
18.4816 | -104.3205
18.4853 | -104.3174
18.4889 | -104.3144
18.4925 | -104.311
18.4961 | -104.3045
18.4998 | -104.2956
18.5034 | -104.2889
18.507 | -104.281
18.5107 | -104.2715
18.5143 | -104.2601
18.5179 | -104.2474
18.5215 | -104.2369
18.5252 | -104.2252
18.5288 | -104.2131
18.5324 | -104.201
18.5361 | -104.1887
18.5397 | -104.178
18.5433 | -104.1638
18.5469 | -104.1465
18.5506 | -104.1273
18.5542 | -104.1059
18.5578 | -104.0867
18.5615 | -104.067
18.5651 | -104.0448
18.5687 | -104.0251
18.5723 | -104.0055
18.576 | -103.9843
18.5796 | -103.961
18.5832 | -103.9393
18.5868 | -103.9163
18.5905 | -103.8915
18.5941 | -103.8689
18.5977 | -103.8476
18.6014 | -103.8235
18.605 | -103.7992
18.6086 | -103.7759
18.6122 | -103.7536
18.6159 | -103.7304
18.6195 | -103.71
18.6231 | -103.6894
18.6268 | -103.6701
18.6304 | -103.6515
18.634 | -103.6321
18.6376 | -103.6134
18.6413 | -103.5964
18.6449 | -103.5804
18.6485 | -103.5639
18.6522 | -103.5507
18.6558 | -103.5396
18.6594 | -103.5274
18.663 | -103.5145
18.6667 | -103.5039
18.6703 | -103.4959
18.6739 | -103.4878
18.6776 | -103.4788
18.6812 | -103.4706
18.6848 | -103.4632
18.6884 | -103.4595
18.6921 | -103.4561
18.6957 | -103.4542
18.6993 | -103.4525
18.7029 | -103.4519
18.7066 | -103.4511
18.7102 | -103.4508
18.7138 | -103.4513
18.7175 | -103.4516
18.7211 | -103.4518
18.7247 | -103.4523
18.7283 | -103.4489
18.732 | -103.4482
18.7356 | -103.4493
18.7392 | -103.4506
18.7429 | -103.4503
18.7465 | -103.4507
18.7501 | -103.4508
18.7537 | -103.4537
18.7574 | -103.4552
18.761 | -103.4579
18.7646 | -103.4604
18.7683 | -103.4595
18.7719 | -103.4574
18.7755 | -103.4569
18.7791 | -103.4574
18.7828 | -103.456
18.7864 | -103.4552
18.79 | -103.4535
18.7937 | -103.4502
18.7973 | -103.4477
18.8009 | -103.4448
18.8045 | -103.4397
18.8082 | -103.4355
18.8118 | -103.4294
18.8154 | -103.4248
18.819 | -103.4203
18.8227 | -103.4112
18.8263 | -103.4045
18.8299 | -103.3976
18.8336 | -103.3898
18.8372 | -103.3815
18.8408 | -103.3726
18.8444 | -103.3631
18.8481 | -103.3507
18.8517 | -103.3369
18.8553 | -103.3229
18.859 | -103.3119
18.8626 | -103.2993
18.8662 | -103.2866
18.8698 | -103.2746
18.8735 | -103.2644
18.8771 | -103.2538
18.8807 | -103.2427
18.8844 | -103.2325
18.888 | -103.2205
18.8916 | -103.2098
18.8952 | -103.1971
18.8989 | -103.1846
18.9025 | -103.1723
18.9061 | -103.159
18.9098 | -103.1461
18.9134 | -103.1332
18.917 | -103.1198
18.9206 | -103.1068
18.9243 | -103.0941
18.9279 | -103.0822
18.9315 | -103.0694
18.9351 | -103.0566
18.9388 | -103.0449
18.9424 | -103.0335
18.946 | -103.0226
18.9497 | -103.0097
18.9533 | -102.999
18.9569 | -102.9879
18.9605 | -102.9757
18.9642 | -102.9636
18.9678 | -102.9552
18.9714 | -102.9464
18.9751 | -102.9381
18.9787 | -102.9313
18.9823 | -102.923
18.9859 | -102.9154
18.9896 | -102.91
18.9932 | -102.9045
18.9968 | -102.9027
19.0005 | -102.9009
19.0041 | -102.8965
19.0077 | -102.8934
19.0113 | -102.8913
19.015 | -102.8877
19.0186 | -102.8836
19.0222 | -102.8775
19.0258 | -102.8757
19.0295 | -102.8729
19.0331 | -102.872
19.0367 | -102.8751
19.0404 | -102.8795
19.044 | -102.883
19.0476 | -102.8865
19.0512 | -102.8905
19.0549 | -102.8929
19.0585 | -102.8963
19.0621 | -102.9007
19.0658 | -102.9034
19.0694 | -102.9043
19.073 | -102.9061
19.0766 | -102.9088
19.0803 | -102.9126
19.0839 | -102.9153
19.0875 | -102.9171
19.0912 | -102.9188
19.0948 | -102.918
19.0984 | -102.9201
19.102 | -102.9228
19.1057 | -102.9245
19.1093 | -102.922
19.1129 | -102.9202
19.1166 | -102.9208
19.1202 | -102.9219
19.1238 | -102.9225
19.1274 | -102.9231
19.1311 | -102.9195
19.1347 | -102.9143
19.1383 | -102.9077
19.142 | -102.9042
19.1456 | -102.9001
19.1492 | -102.8959
19.1528 | -102.8894
19.1565 | -102.8832
19.1601 | -102.8781
19.1637 | -102.8704
19.1673 | -102.8635
19.171 | -102.8565
19.1746 | -102.8525
19.1782 | -102.8481
19.1819 | -102.8426
19.1855 | -102.8365
19.1891 | -102.8318
19.1927 | -102.8268
19.1964 | -102.821
19.2 | -102.815
19.2036 | -102.8069
19.2073 | -102.7999
19.2109 | -102.7954
19.2145 | -102.7914
19.2181 | -102.7875
19.2218 | -102.7845
19.2254 | -102.7799
19.229 | -102.7735
19.2327 | -102.7724
19.2363 | -102.7716
19.2399 | -102.7683
19.2435 | -102.7658
19.2472 | -102.7638
19.2508 | -102.7603
19.2544 | -102.7592
19.2581 | -102.7595
19.2617 | -102.7611
19.2653 | -102.7639
19.2689 | -102.7669
19.2726 | -102.769
19.2762 | -102.7735
19.2798 | -102.78
19.2834 | -102.7869
19.2871 | -102.795
19.2907 | -102.8024
19.2943 | -102.8121
19.298 | -102.8216
19.3016 | -102.8304
19.3052 | -102.8386
19.3088 | -102.8434
19.3125 | -102.849
19.3161 | -102.8576
19.3197 | -102.8669
19.3234 | -102.8737
19.327 | -102.882
19.3306 | -102.8904
19.3342 | -102.8996
19.3379 | -102.9102
19.3415 | -102.9219
19.3451 | -102.9338
19.3488 | -102.9442
19.3524 | -102.9536
19.356 | -102.9623
19.3596 | -102.9717
19.3633 | -102.9788
19.3669 | -102.9885
19.3705 | -102.9997
19.3742 | -103.0068
19.3778 | -103.0138
19.3814 | -103.0219
19.385 | -103.0332
19.3887 | -103.0437
19.3923 | -103.051
19.3959 | -103.0619
19.3995 | -103.0721
19.4032 | -103.0797
19.4068 | -103.0881
19.4104 | -103.0989
19.4141 | -103.1095
19.4177 | -103.116
19.4213 | -103.1198
19.4249 | -103.1239
19.4286 | -103.1334
19.4322 | -103.1426
19.4358 | -103.1498
19.4395 | -103.1569
19.4431 | -103.1647
19.4467 | -103.1724
19.4503 | -103.1793
19.454 | -103.1858
19.4576 | -103.1922
19.4612 | -103.1992
19.4649 | -103.2032
19.4685 | -103.209
19.4721 | -103.2158
19.4757 | -103.2234
19.4794 | -103.2322
19.483 | -103.2394
19.4866 | -103.2419
19.4902 | -103.2465
19.4939 | -103.2518
19.4975 | -103.2588
19.5011 | -103.2666
19.5048 | -103.2757
19.5084 | -103.2799
19.512 | -103.2842
19.5156 | -103.2903
19.5193 | -103.2982
19.5229 | -103.3067
19.5265 | -103.3157
19.5302 | -103.3237
19.5338 | -103.3316
19.5374 | -103.339
19.541 | -103.3478
19.5447 | -103.356
19.5483 | -103.3651
19.5519 | -103.3755
19.5556 | -103.3857
19.5592 | -103.3959
19.5628 | -103.4059
19.5664 | -103.4192
19.5701 | -103.4301
19.5737 | -103.4399
19.5773 | -103.4529
19.581 | -103.4672
19.5846 | -103.4815
19.5882 | -103.4943
19.5918 | -103.5031
19.5955 | -103.5111
19.5991 | -103.5246
19.6027 | -103.5365
19.6063 | -103.5458
19.61 | -103.5563
19.6136 | -103.5656
19.6172 | -103.5738
19.6209 | -103.5807
19.6245 | -103.5896
19.6281 | -103.5983
19.6317 | -103.6071
19.6354 | -103.6142
19.639 | -103.6202
19.6426 | -103.6277
19.6463 | -103.633
19.6499 | -103.6337
19.6535 | -103.6373
19.6571 | -103.6386
19.6608 | -103.6383
19.6644 | -103.6392
19.668 | -103.6387
19.6717 | -103.6391
19.6753 | -103.636
19.6789 | -103.6318
19.6825 | -103.6304
19.6862 | -103.6285
19.6898 | -103.6235
19.6934 | -103.6179
19.6971 | -103.6155
19.7007 | -103.6137
19.7043 | -103.611
19.7079 | -103.606
19.7116 | -103.5995
19.7152 | -103.5961
19.7188 | -103.5941
19.7224 | -103.5913
19.7261 | -103.5892
19.7297 | -103.5854
19.7333 | -103.5833
19.737 | -103.5834
19.7406 | -103.5844
19.7442 | -103.5871
19.7478 | -103.5891
19.7515 | -103.5903
19.7551 | -103.5903
19.7587 | -103.5955
19.7624 | -103.6001
19.766 | -103.6079
19.7696 | -103.6155
19.7732 | -103.6248
19.7769 | -103.6336
19.7805 | -103.6432
19.7841 | -103.6524
19.7878 | -103.6652
19.7914 | -103.6781
19.795 | -103.689
19.7986 | -103.7023
19.8023 | -103.7147
19.8059 | -103.7262
19.8095 | -103.7399
19.8132 | -103.7518
19.8168 | -103.7632
19.8204 | -103.7731
19.824 | -103.7855
19.8277 | -103.798
19.8313 | -103.8075
19.8349 | -103.8166
19.8385 | -103.8287
19.8422 | -103.8394
19.8458 | -103.8475
19.8494 | -103.8571
19.8531 | -103.8651
19.8567 | -103.8713
19.8603 | -103.8767
19.8639 | -103.8839
19.8676 | -103.8917
19.8712 | -103.8996
19.8748 | -103.9081
19.8785 | -103.9123
19.8821 | -103.9162
19.8857 | -103.921
19.8893 | -103.9194
19.893 | -103.9215
19.8966 | -103.9235
19.9002 | -103.9243
19.9039 | -103.9256
19.9075 | -103.9253
19.9111 | -103.9242
19.9147 | -103.9224
19.9184 | -103.9205
19.922 | -103.9193
19.9256 | -103.9171
19.9293 | -103.9155
19.9329 | -103.9138
19.9365 | -103.909
19.9401 | -103.903
19.9438 | -103.8988
19.9474 | -103.8928
19.951 | -103.8879
19.9546 | -103.883
19.9583 | -103.8781
19.9619 | -103.8722
19.9655 | -103.8651
19.9692 | -103.8584
19.9728 | -103.8535
19.9764 | -103.8484
19.98 | -103.8405
19.9837 | -103.8321
19.9873 | -103.8225
19.9909 | -103.8127
19.9946 | -103.8061
19.9982 | -103.797
20.0018 | -103.7881
20.0054 | -103.7806
20.0091 | -103.7712
20.0127 | -103.7627
20.0163 | -103.7563
20.02 | -103.7496
20.0236 | -103.743
20.0272 | -103.7366
20.0308 | -103.7299
20.0345 | -103.7248
20.0381 | -103.7216
20.0417 | -103.7157
20.0454 | -103.7112
20.049 | -103.7055
20.0526 | -103.6982
20.0562 | -103.6927
20.0599 | -103.6888
20.0635 | -103.6843
20.0671 | -103.6822
20.0707 | -103.6787
20.0744 | -103.6755
20.078 | -103.6721
20.0816 | -103.6637
20.0853 | -103.657
20.0889 | -103.6547
20.0925 | -103.6499
20.0961 | -103.644
20.0998 | -103.6369
20.1034 | -103.6292
20.107 | -103.6201
20.1107 | -103.6122
20.1143 | -103.6055
20.1179 | -103.5962
20.1215 | -103.5853
20.1252 | -103.5778
20.1288 | -103.5703
20.1324 | -103.5613
20.1361 | -103.5514
20.1397 | -103.541
20.1433 | -103.5298
20.1469 | -103.5189
20.1506 | -103.5082
20.1542 | -103.497
20.1578 | -103.4846
20.1615 | -103.4721
20.1651 | -103.4588
20.1687 | -103.4478
20.1723 | -103.4375
20.176 | -103.4272
20.1796 | -103.4182
20.1832 | -103.411
20.1868 | -103.405
20.1905 | -103.3972
20.1941 | -103.3896
20.1977 | -103.3841
20.2014 | -103.3785
20.205 | -103.3763
20.2086 | -103.3742
20.2122 | -103.3728
20.2159 | -103.3719
20.2195 | -103.3734
20.2231 | -103.3766
20.2268 | -103.3811
20.2304 | -103.3841
20.234 | -103.3869
20.2376 | -103.3892
20.2413 | -103.392
20.2449 | -103.3954
20.2485 | -103.402
20.2522 | -103.4093
20.2558 | -103.416
20.2594 | -103.4188
20.263 | -103.4224
20.2667 | -103.43
20.2703 | -103.4375
20.2739 | -103.4435
20.2776 | -103.4512
20.2812 | -103.4569
20.2848 | -103.4638
20.2884 | -103.4706
20.2921 | -103.4756
20.2957 | -103.4795
20.2993 | -103.4832
20.3029 | -103.4854
20.3066 | -103.4867
20.3102 | -103.4872
20.3138 | -103.489
20.3175 | -103.4903
20.3211 | -103.491
20.3247 | -103.4894
20.3283 | -103.4852
20.332 | -103.4816
20.3356 | -103.4763
20.3392 | -103.4711
20.3429 | -103.4626
20.3465 | -103.4538
20.3501 | -103.4438
20.3537 | -103.4334
20.3574 | -103.4243
20.361 | -103.4166
20.3646 | -103.4074
20.3683 | -103.3986
20.3719 | -103.3891
20.3755 | -103.3766
20.3791 | -103.3651
20.3828 | -103.3549
20.3864 | -103.3443
20.39 | -103.3327
20.3937 | -103.3241
20.3973 | -103.3123
20.4009 | -103.2998
20.4045 | -103.2903
20.4082 | -103.2808
20.4118 | -103.2762
20.4154 | -103.2659
20.419 | -103.2568
20.4227 | -103.2497
20.4263 | -103.2428
20.4299 | -103.2364
20.4336 | -103.2312
20.4372 | -103.2246
20.4408 | -103.2187
20.4444 | -103.2154
20.4481 | -103.2106
20.4517 | -103.2089
20.4553 | -103.2076
20.459 | -103.204
20.4626 | -103.2017
20.4662 | -103.2007
20.4698 | -103.1988
20.4735 | -103.1968
20.4771 | -103.1952
20.4807 | -103.1957
20.4844 | -103.1975
20.488 | -103.1976
20.4916 | -103.198
20.4952 | -103.1983
20.4989 | -103.2002
20.5025 | -103.2003
20.5061 | -103.201
20.5098 | -103.202
20.5134 | -103.2043
20.517 | -103.2055
20.5206 | -103.2067
20.5243 | -103.2098
20.5279 | -103.2094
20.5315 | -103.2083
20.5351 | -103.21
20.5388 | -103.2115
20.5424 | -103.213
20.546 | -103.2147
20.5497 | -103.2172
20.5533 | -103.2173
20.5569 | -103.2143
20.5605 | -103.212
20.5642 | -103.212
20.5678 | -103.2099
20.5714 | -103.209
20.5751 | -103.2088
20.5787 | -103.208
20.5823 | -103.207
20.5859 | -103.2084
20.5896 | -103.2068
20.5932 | -103.204
20.5968 | -103.203
20.6005 | -103.202
20.6041 | -103.1997
20.6077 | -103.1974
20.6113 | -103.1938
20.615 | -103.1948
20.6186 | -103.1957
20.6222 | -103.1955
20.6258 | -103.1952
20.6295 | -103.1946
20.6331 | -103.1939
20.6367 | -103.1915
20.6404 | -103.1907
20.644 | -103.1885
20.6476 | -103.1859
20.6512 | -103.1836
20.6549 | -103.1833
20.6585 | -103.1829
20.6621 | -103.181
20.6658 | -103.1779
20.6694 | -103.1748
20.673 | -103.1719
20.6766 | -103.1693
20.6803 | -103.1682
20.6839 | -103.1654
20.6875 | -103.1653
20.6912 | -103.166
20.6948 | -103.1654
20.6984 | -103.1645
20.702 | -103.1647
20.7057 | -103.1652
20.7093 | -103.1649
20.7129 | -103.1636
20.7166 | -103.1631
20.7202 | -103.1653
20.7238 | -103.1694
20.7274 | -103.1712
20.7311 | -103.1746
20.7347 | -103.1788
20.7383 | -103.1822
20.7419 | -103.1876
20.7456 | -103.1927
20.7492 | -103.1971
20.7528 | -103.2012
20.7565 | -103.2024
20.7601 | -103.2066
20.7637 | -103.211
20.7673 | -103.2146
20.771 | -103.2186
20.7746 | -103.2233
20.7782 | -103.2256
20.7819 | -103.2278
20.7855 | -103.2324
20.7891 | -103.2373
20.7927 | -103.2408
20.7964 | -103.2449
20.8 | -103.2493
20.8036 | -103.254
20.8073 | -103.255
20.8109 | -103.2565
20.8145 | -103.2598
20.8181 | -103.2624
20.8218 | -103.2671
20.8254 | -103.2704
20.829 | -103.2688
20.8327 | -103.2682
20.8363 | -103.2666
20.8399 | -103.2647
20.8435 | -103.2651
20.8472 | -103.2665
20.8508 | -103.2666
20.8544 | -103.2662
20.858 | -103.2653
20.8617 | -103.2639
20.8653 | -103.2601
20.8689 | -103.2546
20.8726 | -103.2517
20.8762 | -103.2501
20.8798 | -103.2474
20.8834 | -103.2468
20.8871 | -103.2474
20.8907 | -103.2471
20.8943 | -103.2472
20.898 | -103.2463
20.9016 | -103.2444
20.9052 | -103.244
20.9088 | -103.2396
20.9125 | -103.2376
20.9161 | -103.2378
20.9197 | -103.2405
20.9234 | -103.244
20.927 | -103.2478
20.9306 | -103.2525
20.9342 | -103.2592
20.9379 | -103.2648
20.9415 | -103.2713
20.9451 | -103.2774
20.9488 | -103.2834
20.9524 | -103.29
20.956 | -103.2981
20.9596 | -103.3059
20.9633 | -103.3159
20.9669 | -103.3262
20.9705 | -103.3358
20.9742 | -103.345
20.9778 | -103.3546
20.9814 | -103.3643
20.985 | -103.3735
20.9887 | -103.3846
20.9923 | -103.3955
20.9959 | -103.4052
20.9995 | -103.4154
21.0032 | -103.426
21.0068 | -103.4375
21.0104 | -103.4508
21.0141 | -103.4623
21.0177 | -103.4717
21.0213 | -103.4796
21.0249 | -103.4907
21.0286 | -103.4987
21.0322 | -103.5071
21.0358 | -103.5168
21.0395 | -103.5262
21.0431 | -103.5338
21.0467 | -103.5426
21.0503 | -103.553
21.054 | -103.5657
21.0576 | -103.5755
21.0612 | -103.5864
21.0649 | -103.5974
21.0685 | -103.6081
21.0721 | -103.6171
21.0757 | -103.6284
21.0794 | -103.6413
21.083 | -103.6505
21.0866 | -103.6564
21.0902 | -103.665
21.0939 | -103.6759
21.0975 | -103.686
21.1011 | -103.6949
21.1048 | -103.7057
21.1084 | -103.7164
21.112 | -103.724
21.1156 | -103.7321
21.1193 | -103.7412
21.1229 | -103.7501
21.1265 | -103.7598
21.1302 | -103.766
21.1338 | -103.7761
21.1374 | -103.7857
21.141 | -103.7954
21.1447 | -103.805
21.1483 | -103.8147
21.1519 | -103.8222
21.1556 | -103.8285
21.1592 | -103.8332
21.1628 | -103.8405
21.1664 | -103.8501
21.1701 | -103.8573
21.1737 | -103.8623
21.1773 | -103.8672
21.181 | -103.8734
21.1846 | -103.8786
21.1882 | -103.8822
21.1918 | -103.884
21.1955 | -103.8885
21.1991 | -103.8929
21.2027 | -103.8966
21.2063 | -103.8996
21.21 | -103.9015
21.2136 | -103.9021
21.2172 | -103.9023
21.2209 | -103.9018
21.2245 | -103.9008
21.2281 | -103.9015
21.2317 | -103.9035
21.2354 | -103.902
21.239 | -103.8997
21.2426 | -103.8976
21.2463 | -103.895
21.2499 | -103.8921
21.2535 | -103.8893
21.2571 | -103.8833
21.2608 | -103.8802
21.2644 | -103.8762
21.268 | -103.8726
21.2717 | -103.871
21.2753 | -103.869
21.2789 | -103.8663
21.2825 | -103.8635
21.2862 | -103.8617
21.2898 | -103.8581
21.2934 | -103.8556
21.2971 | -103.8539
21.3007 | -103.8521
21.3043 | -103.8509
21.3079 | -103.8496
21.3116 | -103.8453
21.3152 | -103.8432
21.3188 | -103.8426
21.3224 | -103.8439
21.3261 | -103.8448
21.3297 | -103.8486
21.3333 | -103.8477
21.337 | -103.8454
21.3406 | -103.8438
21.3442 | -103.8388
21.3478 | -103.8365
21.3515 | -103.8347
21.3551 | -103.8303
21.3587 | -103.8271
21.3624 | -103.8227
21.366 | -103.8211
21.3696 | -103.8197
21.3732 | -103.8176
21.3769 | -103.8157
21.3805 | -103.8133
21.3841 | -103.8101
21.3878 | -103.8083
21.3914 | -103.8053
21.395 | -103.7993
21.3986 | -103.7935
21.4023 | -103.7874
21.4059 | -103.7823
21.4095 | -103.7767
21.4132 | -103.7712
21.4168 | -103.7603
21.4204 | -103.7478
21.424 | -103.7373
21.4277 | -103.7281
21.4313 | -103.7185
21.4349 | -103.7042
21.4385 | -103.6896
21.4422 | -103.6738
21.4458 | -103.6592
21.4494 | -103.6431
21.4531 | -103.6297
21.4567 | -103.6166
21.4603 | -103.5993
21.4639 | -103.5873
21.4676 | -103.5744
21.4712 | -103.5579
21.4748 | -103.5396
21.4785 | -103.5223
21.4821 | -103.5052
21.4857 | -103.4871
21.4893 | -103.4698
21.493 | -103.4522
21.4966 | -103.4356
21.5002 | -103.418
21.5039 | -103.3977
21.5075 | -103.3787
21.5111 | -103.3618
21.5147 | -103.3428
21.5184 | -103.3247
21.522 | -103.3054
21.5256 | -103.2863
21.5293 | -103.2689
21.5329 | -103.2494
21.5365 | -103.2285
21.5401 | -103.2086
21.5438 | -103.1898
21.5474 | -103.1694
21.551 | -103.15
21.5546 | -103.1311
21.5583 | -103.1115
21.5619 | -103.096
21.5655 | -103.0777
21.5692 | -103.0591
21.5728 | -103.0397
21.5764 | -103.0196
21.58 | -103.0009
21.5837 | -102.9813
21.5873 | -102.9616
21.5909 | -102.9446
21.5946 | -102.9283
21.5982 | -102.9096
21.6018 | -102.8933
21.6054 | -102.8772
21.6091 | -102.8597
21.6127 | -102.8423
21.6163 | -102.8284
21.62 | -102.813
21.6236 | -102.7969
21.6272 | -102.7805
21.6308 | -102.7641
21.6345 | -102.7487
21.6381 | -102.7336
21.6417 | -102.7185
21.6454 | -102.7035
21.649 | -102.6925
21.6526 | -102.68
21.6562 | -102.6669
21.6599 | -102.6553
21.6635 | -102.6419
21.6671 | -102.6296
21.6707 | -102.6178
21.6744 | -102.6072
21.678 | -102.5961
21.6816 | -102.585
21.6853 | -102.5739
21.6889 | -102.5653
21.6925 | -102.5574
21.6961 | -102.5502
21.6998 | -102.5436
21.7034 | -102.5373
21.707 | -102.5318
21.7107 | -102.5273
21.7143 | -102.5235
21.7179 | -102.5198
21.7215 | -102.5198
21.7252 | -102.5203
21.7288 | -102.5169
21.7324 | -102.5124
21.7361 | -102.5085
21.7397 | -102.5078
21.7433 | -102.5088
21.7469 | -102.5097
21.7506 | -102.5082
21.7542 | -102.508
21.7578 | -102.508
21.7615 | -102.5089
21.7651 | -102.5107
21.7687 | -102.5147
21.7723 | -102.5178
21.776 | -102.5195
21.7796 | -102.5232
21.7832 | -102.5283
21.7868 | -102.5328
21.7905 | -102.5392
21.7941 | -102.5441
21.7977 | -102.5477
21.8014 | -102.552
21.805 | -102.5567
21.8086 | -102.5608
21.8122 | -102.5665
21.8159 | -102.5689
21.8195 | -102.573
21.8231 | -102.578
21.8268 | -102.5785
21.8304 | -102.5784
21.834 | -102.5791
21.8376 | -102.5787
21.8413 | -102.5795
21.8449 | -102.5822
21.8485 | -102.5849
21.8522 | -102.5849
21.8558 | -102.5832
21.8594 | -102.5825
21.863 | -102.582
21.8667 | -102.5789
21.8703 | -102.5725
21.8739 | -102.5673
21.8776 | -102.562
21.8812 | -102.5569
21.8848 | -102.5517
21.8884 | -102.5447
21.8921 | -102.5384
21.8957 | -102.5308
21.8993 | -102.5198
21.9029 | -102.5082
21.9066 | -102.4988
21.9102 | -102.4886
21.9138 | -102.4782
21.9175 | -102.4698
21.9211 | -102.4618
21.9247 | -102.4519
21.9283 | -102.4406
21.932 | -102.4314
21.9356 | -102.4227
21.9392 | -102.414
21.9429 | -102.4066
21.9465 | -102.399
21.9501 | -102.3916
21.9537 | -102.3847
21.9574 | -102.376
21.961 | -102.3668
21.9646 | -102.3587
21.9683 | -102.3502
21.9719 | -102.3426
21.9755 | -102.3355
21.9791 | -102.3296
21.9828 | -102.3249
21.9864 | -102.318
21.99 | -102.3131
21.9937 | -102.3075
21.9973 | -102.3025
22.0009 | -102.2967
22.0045 | -102.2888
22.0082 | -102.2797
22.0118 | -102.2735
22.0154 | -102.2692
22.019 | -102.2655
22.0227 | -102.2596
22.0263 | -102.2527
22.0299 | -102.2466
22.0336 | -102.2388
22.0372 | -102.2313
22.0408 | -102.2272
22.0444 | -102.2239
22.0481 | -102.2195
22.0517 | -102.2134
22.0553 | -102.206
22.059 | -102.1981
22.0626 | -102.1927
22.0662 | -102.1868
22.0698 | -102.1805
22.0735 | -102.174
22.0771 | -102.1658
22.0807 | -102.1567
22.0844 | -102.149
22.088 | -102.1419
22.0916 | -102.1347
22.0952 | -102.1267
22.0989 | -102.1219
22.1025 | -102.114
22.1061 | -102.1068
22.1098 | -102.1008
22.1134 | -102.0921
22.117 | -102.0846
22.1206 | -102.08
22.1243 | -102.076
22.1279 | -102.0717
22.1315 | -102.0649
22.1351 | -102.0575
22.1388 | -102.053
22.1424 | -102.0479
22.146 | -102.0417
22.1497 | -102.0348
22.1533 | -102.0247
22.1569 | -102.015
22.1605 | -102.007
22.1642 | -101.9998
22.1678 | -101.994
22.1714 | -101.9862
22.1751 | -101.9779
22.1787 | -101.9704
22.1823 | -101.9637
22.1859 | -101.9569
22.1896 | -101.9495
22.1932 | -101.943
22.1968 | -101.936
22.2005 | -101.9293
22.2041 | -101.9224
22.2077 | -101.915
22.2113 | -101.9102
22.215 | -101.9056
22.2186 | -101.8996
22.2222 | -101.8937
22.2259 | -101.8877
22.2295 | -101.8794
22.2331 | -101.8718
22.2367 | -101.8628
22.2404 | -101.8553
22.244 | -101.8488
22.2476 | -101.8426
22.2512 | -101.8381
22.2549 | -101.8313
22.2585 | -101.8242
22.2621 | -101.8179
22.2658 | -101.813
22.2694 | -101.8073
22.273 | -101.802
22.2766 | -101.7968
22.2803 | -101.7923
22.2839 | -101.7882
22.2875 | -101.7822
22.2912 | -101.7744
22.2948 | -101.7693
22.2984 | -101.7647
22.302 | -101.7573
22.3057 | -101.7491
22.3093 | -101.7433
22.3129 | -101.7378
22.3166 | -101.7327
22.3202 | -101.728
22.3238 | -101.7235
22.3274 | -101.7162
22.3311 | -101.7079
22.3347 | -101.7003
22.3383 | -101.6941
22.342 | -101.6895
22.3456 | -101.6852
22.3492 | -101.6797
22.3528 | -101.6725
22.3565 | -101.6652
22.3601 | -101.6578
22.3637 | -101.6514
22.3673 | -101.6464
22.371 | -101.6425
22.3746 | -101.6381
22.3782 | -101.6306
22.3819 | -101.6242
22.3855 | -101.6174
22.3891 | -101.6096
22.3927 | -101.6032
22.3964 | -101.5982
22.4 | -101.5918
22.4036 | -101.5862
22.4073 | -101.5818
22.4109 | -101.5775
22.4145 | -101.5717
22.4181 | -101.5672
22.4218 | -101.5626
22.4254 | -101.5588
22.429 | -101.5576
22.4327 | -101.5537
22.4363 | -101.5482
22.4399 | -101.5422
22.4435 | -101.5375
22.4472 | -101.5312
22.4508 | -101.526
22.4544 | -101.5214
22.4581 | -101.5173
22.4617 | -101.5144
22.4653 | -101.5106
22.4689 | -101.5073
22.4726 | -101.5025
22.4762 | -101.4957
22.4798 | -101.4882
22.4834 | -101.481
22.4871 | -101.4747
22.4907 | -101.4668
22.4943 | -101.4576
22.498 | -101.4486
22.5016 | -101.4379
22.5052 | -101.4197
22.5088 | -101.4083
22.5125 | -101.3946
22.5161 | -101.379
22.5197 | -101.3639
22.5234 | -101.3484
22.527 | -101.3301
22.5306 | -101.3083
22.5342 | -101.2887
22.5379 | -101.2722
22.5415 | -101.2549
22.5451 | -101.2352
22.5488 | -101.2167
22.5524 | -101.1967
22.556 | -101.1762
22.5596 | -101.1541
22.5633 | -101.13
22.5669 | -101.107
22.5705 | -101.0842
22.5741 | -101.0617
22.5778 | -101.0395
22.5814 | -101.018
22.585 | -100.9957
22.5887 | -100.9746
22.5923 | -100.9531
22.5959 | -100.9321
22.5995 | -100.9129
22.6032 | -100.8932
22.6068 | -100.8753
22.6104 | -100.8593
22.6141 | -100.8446
22.6177 | -100.8312
22.6213 | -100.8205
22.6249 | -100.8112
22.6286 | -100.8006
22.6322 | -100.7906
22.6358 | -100.7819
22.6395 | -100.7756
22.6431 | -100.7705
22.6467 | -100.7661
22.6503 | -100.7621
22.654 | -100.7601
22.6576 | -100.7592
22.6612 | -100.7606
22.6649 | -100.7632
22.6685 | -100.7686
22.6721 | -100.7762
22.6757 | -100.784
22.6794 | -100.794
22.683 | -100.8031
22.6866 | -100.8123
22.6902 | -100.8234
22.6939 | -100.8362
22.6975 | -100.8502
22.7011 | -100.8651
22.7048 | -100.881
22.7084 | -100.8982
22.712 | -100.9146
22.7156 | -100.931
22.7193 | -100.948
22.7229 | -100.9643
22.7265 | -100.9813
22.7302 | -100.9987
22.7338 | -101.0154
22.7374 | -101.0331
22.741 | -101.0516
22.7447 | -101.0672
22.7483 | -101.0805
22.7519 | -101.0964
22.7556 | -101.111
22.7592 | -101.1239
22.7628 | -101.1361
22.7664 | -101.1494
22.7701 | -101.1614
22.7737 | -101.1715
22.7773 | -101.1811
22.781 | -101.1897
22.7846 | -101.1979
22.7882 | -101.2059
22.7918 | -101.2121
22.7955 | -101.2154
22.7991 | -101.2183
22.8027 | -101.2191
22.8063 | -101.2174
22.81 | -101.2166
22.8136 | -101.2165
22.8172 | -101.2127
22.8209 | -101.2068
22.8245 | -101.2036
22.8281 | -101.1984
22.8317 | -101.1915
22.8354 | -101.1844
22.839 | -101.1789
22.8426 | -101.1735
22.8463 | -101.1674
22.8499 | -101.1626
22.8535 | -101.1577
22.8571 | -101.1482
22.8608 | -101.1335
22.8644 | -101.1186
22.868 | -101.1035
22.8717 | -101.0911
22.8753 | -101.0803
22.8789 | -101.0695
22.8825 | -101.0605
22.8862 | -101.0516
22.8898 | -101.0419
22.8934 | -101.0306
22.8971 | -101.0206
22.9007 | -101.0126
22.9043 | -101.0025
22.9079 | -100.9943
22.9116 | -100.9866
22.9152 | -100.9787
22.9188 | -100.9713
22.9224 | -100.9608
22.9261 | -100.952
22.9297 | -100.9462
22.9333 | -100.9392
22.937 | -100.9298
22.9406 | -100.9218
22.9442 | -100.9148
22.9478 | -100.909
22.9515 | -100.9041
22.9551 | -100.8991
22.9587 | -100.8954
22.9624 | -100.89
22.966 | -100.8846
22.9696 | -100.8812
22.9732 | -100.8804
22.9769 | -100.8764
22.9805 | -100.8726
22.9841 | -100.8686
22.9878 | -100.8657
22.9914 | -100.8637
22.995 | -100.8626
22.9986 | -100.8619
23.0023 | -100.8599
23.0059 | -100.8605
23.0095 | -100.8615
23.0132 | -100.8635
23.0168 | -100.8644
23.0204 | -100.8671
23.024 | -100.8692
23.0277 | -100.8784
23.0313 | -100.883
23.0349 | -100.8876
23.0385 | -100.8916
23.0422 | -100.8954
23.0458 | -100.8992
23.0494 | -100.9041
23.0531 | -100.908
23.0567 | -100.9133
23.0603 | -100.9187
23.0639 | -100.9225
23.0676 | -100.9263
23.0712 | -100.932
23.0748 | -100.9397
23.0785 | -100.9452
23.0821 | -100.9512
23.0857 | -100.9581
23.0893 | -100.9618
23.093 | -100.9649
23.0966 | -100.9676
23.1002 | -100.97
23.1039 | -100.9736
23.1075 | -100.9759
23.1111 | -100.9782
23.1147 | -100.9792
23.1184 | -100.9801
23.122 | -100.9843
23.1256 | -100.9874
23.1293 | -100.9909
23.1329 | -100.9922
23.1365 | -100.9898
23.1401 | -100.9875
23.1438 | -100.9857
23.1474 | -100.9864
23.151 | -100.9818
23.1546 | -100.9776
23.1583 | -100.9738
23.1619 | -100.9704
23.1655 | -100.9664
23.1692 | -100.9616
23.1728 | -100.9564
23.1764 | -100.9524
23.18 | -100.9482
23.1837 | -100.9454
23.1873 | -100.945
23.1909 | -100.9425
23.1946 | -100.9411
23.1982 | -100.9376
23.2018 | -100.9328
23.2054 | -100.9295
23.2091 | -100.9258
23.2127 | -100.921
23.2163 | -100.9165
23.22 | -100.9147
23.2236 | -100.9144
23.2272 | -100.9143
23.2308 | -100.914
23.2345 | -100.9118
23.2381 | -100.9123
23.2417 | -100.9146
23.2454 | -100.9163
23.249 | -100.9206
23.2526 | -100.9264
23.2562 | -100.9306
23.2599 | -100.9346
23.2635 | -100.9402
23.2671 | -100.9479
23.2707 | -100.9562
23.2744 | -100.9628
23.278 | -100.9697
23.2816 | -100.9784
23.2853 | -100.9877
23.2889 | -100.9994
23.2925 | -101.0115
23.2961 | -101.0235
23.2998 | -101.0352
23.3034 | -101.0481
23.307 | -101.0623
23.3107 | -101.076
23.3143 | -101.0893
23.3179 | -101.1026
23.3215 | -101.1157
23.3252 | -101.1279
23.3288 | -101.1399
23.3324 | -101.1537
23.3361 | -101.1681
23.3397 | -101.1808
23.3433 | -101.1953
23.3469 | -101.2087
23.3506 | -101.221
23.3542 | -101.2309
23.3578 | -101.2399
23.3615 | -101.2496
23.3651 | -101.2601
23.3687 | -101.2725
23.3723 | -101.2843
23.376 | -101.297
23.3796 | -101.3073
23.3832 | -101.3168
23.3868 | -101.3254
23.3905 | -101.3342
23.3941 | -101.3427
23.3977 | -101.3504
23.4014 | -101.3586
23.405 | -101.3648
23.4086 | -101.37
23.4122 | -101.3749
23.4159 | -101.38
23.4195 | -101.3878
23.4231 | -101.3935
23.4268 | -101.3995
23.4304 | -101.4054
23.434 | -101.4087
23.4376 | -101.4117
23.4413 | -101.4155
23.4449 | -101.4212
23.4485 | -101.4279
23.4522 | -101.4326
23.4558 | -101.4357
23.4594 | -101.4386
23.463 | -101.4449
23.4667 | -101.4534
23.4703 | -101.4592
23.4739 | -101.4663
23.4776 | -101.4726
23.4812 | -101.481
23.4848 | -101.4876
23.4884 | -101.4945
23.4921 | -101.5025
23.4957 | -101.5096
23.4993 | -101.5168
23.5029 | -101.5228
23.5066 | -101.5293
23.5102 | -101.5371
23.5138 | -101.5442
23.5175 | -101.5523
23.5211 | -101.564
23.5247 | -101.5756
23.5283 | -101.5868
23.532 | -101.5982
23.5356 | -101.6074
23.5392 | -101.6172
23.5429 | -101.6302
23.5465 | -101.6416
23.5501 | -101.6518
23.5537 | -101.6636
23.5574 | -101.672
23.561 | -101.6817
23.5646 | -101.6941
23.5683 | -101.707
23.5719 | -101.7174
23.5755 | -101.7297
23.5791 | -101.743
23.5828 | -101.7554
23.5864 | -101.7676
23.59 | -101.78
23.5937 | -101.7948
23.5973 | -101.8101
23.6009 | -101.8215
23.6045 | -101.8327
23.6082 | -101.8467
23.6118 | -101.8609
23.6154 | -101.8744
23.619 | -101.8846
23.6227 | -101.8954
23.6263 | -101.9064
23.6299 | -101.9168
23.6336 | -101.9254
23.6372 | -101.9342
23.6408 | -101.9424
23.6444 | -101.95
23.6481 | -101.9594
23.6517 | -101.9686
23.6553 | -101.9777
23.659 | -101.9869
23.6626 | -101.9971
23.6662 | -102.0058
23.6698 | -102.0125
23.6735 | -102.0188
23.6771 | -102.0257
23.6807 | -102.0335
23.6844 | -102.0414
23.688 | -102.0468
23.6916 | -102.0537
23.6952 | -102.0619
23.6989 | -102.0691
23.7025 | -102.0735
23.7061 | -102.078
23.7098 | -102.0836
23.7134 | -102.0909
23.717 | -102.0988
23.7206 | -102.1047
23.7243 | -102.1103
23.7279 | -102.1184
23.7315 | -102.1271
23.7351 | -102.1348
23.7388 | -102.1436
23.7424 | -102.1534
23.746 | -102.1623
23.7497 | -102.1691
23.7533 | -102.1752
23.7569 | -102.1836
23.7605 | -102.1901
23.7642 | -102.1962
23.7678 | -102.2033
23.7714 | -102.2114
23.7751 | -102.2181
23.7787 | -102.2233
23.7823 | -102.2295
23.7859 | -102.2357
23.7896 | -102.2441
23.7932 | -102.2539
23.7968 | -102.2631
23.8005 | -102.272
23.8041 | -102.2797
23.8077 | -102.2877
23.8113 | -102.2938
23.815 | -102.2991
23.8186 | -102.3038
23.8222 | -102.3118
23.8258 | -102.3186
23.8295 | -102.3264
23.8331 | -102.3331
23.8367 | -102.3432
23.8404 | -102.3519
23.844 | -102.3598
23.8476 | -102.3689
23.8512 | -102.376
23.8549 | -102.3812
23.8585 | -102.3869
23.8621 | -102.3917
23.8658 | -102.3954
23.8694 | -102.4004
23.873 | -102.405
23.8766 | -102.4086
23.8803 | -102.4148
23.8839 | -102.4197
23.8875 | -102.4236
23.8912 | -102.427
23.8948 | -102.4333
23.8984 | -102.4369
23.902 | -102.4395
23.9057 | -102.4417
23.9093 | -102.4446
23.9129 | -102.4451
23.9166 | -102.4462
23.9202 | -102.4485
23.9238 | -102.4499
23.9274 | -102.4512
23.9311 | -102.4511
23.9347 | -102.4491
23.9383 | -102.4486
23.9419 | -102.4476
23.9456 | -102.4474
23.9492 | -102.4458
23.9528 | -102.4423
23.9565 | -102.4419
23.9601 | -102.4401
23.9637 | -102.4381
23.9673 | -102.4364
23.971 | -102.4336
23.9746 | -102.4316
23.9782 | -102.4286
23.9819 | -102.4236
23.9855 | -102.4149
23.9891 | -102.409
23.9927 | -102.4024
23.9964 | -102.3966
24 | -102.3908
24.0036 | -102.3838
24.0073 | -102.3766
24.0109 | -102.3706
24.0145 | -102.3634
24.0181 | -102.3556
24.0218 | -102.3492
24.0254 | -102.3453
24.029 | -102.3377
24.0327 | -102.3259
24.0363 | -102.317
24.0399 | -102.3112
24.0435 | -102.3045
24.0472 | -102.2973
24.0508 | -102.2913
24.0544 | -102.2841
24.0581 | -102.2772
24.0617 | -102.2718
24.0653 | -102.2655
24.0689 | -102.2581
24.0726 | -102.2508
24.0762 | -102.2485
24.0798 | -102.2425
24.0834 | -102.2382
24.0871 | -102.2349
24.0907 | -102.2341
24.0943 | -102.2342
24.098 | -102.2317
24.1016 | -102.2299
24.1052 | -102.2287
24.1088 | -102.2282
24.1125 | -102.2271
24.1161 | -102.2273
24.1197 | -102.2277
24.1234 | -102.2271
24.127 | -102.2285
24.1306 | -102.2288
24.1342 | -102.2278
24.1379 | -102.2279
24.1415 | -102.2299
24.1451 | -102.2342
24.1488 | -102.2374
24.1524 | -102.242
24.156 | -102.2466
24.1596 | -102.252
24.1633 | -102.2566
24.1669 | -102.2618
24.1705 | -102.2675
24.1742 | -102.2736
24.1778 | -102.2784
24.1814 | -102.2845
24.185 | -102.29
24.1887 | -102.2946
24.1923 | -102.299
24.1959 | -102.3047
24.1995 | -102.3108
24.2032 | -102.3186
24.2068 | -102.3278
24.2104 | -102.3364
24.2141 | -102.3439
24.2177 | -102.3486
24.2213 | -102.3544
24.2249 | -102.3607
24.2286 | -102.3674
24.2322 | -102.3751
24.2358 | -102.3785
24.2395 | -102.3819
24.2431 | -102.385
24.2467 | -102.388
24.2503 | -102.3897
24.254 | -102.3911
24.2576 | -102.393
24.2612 | -102.3945
24.2649 | -102.3937
24.2685 | -102.3908
24.2721 | -102.389
24.2757 | -102.3896
24.2794 | -102.3896
24.283 | -102.3876
24.2866 | -102.3846
24.2902 | -102.3828
24.2939 | -102.3829
24.2975 | -102.3812
24.3011 | -102.3788
24.3048 | -102.3778
24.3084 | -102.3763
24.312 | -102.3751
24.3156 | -102.3728
24.3193 | -102.3711
24.3229 | -102.3709
24.3265 | -102.3697
24.3302 | -102.3665
24.3338 | -102.3646
24.3374 | -102.3637
24.341 | -102.3626
24.3447 | -102.3624
24.3483 | -102.3613
24.3519 | -102.3599
24.3556 | -102.3584
24.3592 | -102.3574
24.3628 | -102.3553
24.3664 | -102.3528
24.3701 | -102.3517
24.3737 | -102.3526
24.3773 | -102.3529
24.381 | -102.3531
24.3846 | -102.3539
24.3882 | -102.3532
24.3918 | -102.3522
24.3955 | -102.3533
24.3991 | -102.3557
24.4027 | -102.3577
24.4063 | -102.3592
24.41 | -102.3621
24.4136 | -102.3656
24.4172 | -102.3674
24.4209 | -102.3684
24.4245 | -102.3711
24.4281 | -102.3732
24.4317 | -102.3761
24.4354 | -102.3787
24.439 | -102.3798
24.4426 | -102.383
24.4463 | -102.3871
24.4499 | -102.3922
24.4535 | -102.3971
24.4571 | -102.4027
24.4608 | -102.4094
24.4644 | -102.4168
24.468 | -102.4236
24.4717 | -102.4279
24.4753 | -102.4331
24.4789 | -102.4409
24.4825 | -102.4498
24.4862 | -102.4575
24.4898 | -102.4666
24.4934 | -102.4755
24.4971 | -102.4833
24.5007 | -102.4936
24.5043 | -102.5046
24.5079 | -102.5136
24.5116 | -102.5216
24.5152 | -102.531
24.5188 | -102.5426
24.5224 | -102.5541
24.5261 | -102.5639
24.5297 | -102.577
24.5333 | -102.5906
24.537 | -102.6038
24.5406 | -102.616
24.5442 | -102.6273
24.5478 | -102.6384
24.5515 | -102.6498
24.5551 | -102.6604
24.5587 | -102.6688
24.5624 | -102.675
24.566 | -102.6787
24.5696 | -102.6857
24.5732 | -102.6915
24.5769 | -102.6937
24.5805 | -102.6953
24.5841 | -102.6952
24.5878 | -102.6883
24.5914 | -102.6759
24.595 | -102.6405
24.5986 | -102.5394
24.6023 | -102.5021
24.6059 | -102.4679
24.6095 | -102.4367
24.6132 | -102.4033
24.6168 | -102.3675
24.6204 | -102.3257
24.624 | -102.2814
24.6277 | -102.2357
24.6313 | -102.1883
24.6349 | -102.1374
24.6385 | -102.0843
24.6422 | -102.0276
24.6458 | -101.9681
24.6494 | -101.9059
24.6531 | -101.8393
24.6567 | -101.7705
24.6603 | -101.7015
24.6639 | -101.6328
24.6676 | -101.5615
24.6712 | -101.4885
24.6748 | -101.4169
24.6785 | -101.3433
24.6821 | -101.27
24.6857 | -101.1966
24.6893 | -101.1217
24.693 | -101.0479
24.6966 | -100.9771
24.7002 | -100.9074
24.7039 | -100.8404
24.7075 | -100.7753
24.7111 | -100.7118
24.7147 | -100.6517
24.7184 | -100.5935
24.722 | -100.539
24.7256 | -100.4871
24.7293 | -100.4384
24.7329 | -100.392
24.7365 | -100.3502
24.7401 | -100.3133
24.7438 | -100.2801
24.7474 | -100.2513
24.751 | -100.2275
24.7546 | -100.2089
24.7583 | -100.1941
24.7619 | -100.1833
24.7655 | -100.1786
24.7692 | -100.1796
24.7728 | -100.1845
24.7764 | -100.194
24.78 | -100.2076
24.7837 | -100.2259
24.7873 | -100.2499
24.7909 | -100.2779
24.7946 | -100.3098
24.7982 | -100.3451
24.8018 | -100.3852
24.8054 | -100.4292
24.8091 | -100.4766
24.8127 | -100.5282
24.8163 | -100.5833
24.82 | -100.6437
24.8236 | -100.708
24.8272 | -100.7731
24.8308 | -100.8401
24.8345 | -100.9096
24.8381 | -100.9811
24.8417 | -101.0549
24.8454 | -101.1312
24.849 | -101.2078
24.8526 | -101.2845
24.8562 | -101.3626
24.8599 | -101.4417
24.8635 | -101.5205
24.8671 | -101.5988
24.8707 | -101.6769
24.8744 | -101.7552
24.878 | -101.8304
24.8816 | -101.9046
24.8853 | -101.9781
24.8889 | -102.0489
24.8925 | -102.1182
24.8961 | -102.1851
24.8998 | -102.2493
24.9034 | -102.313
24.907 | -102.3755
24.9107 | -102.4372
24.9143 | -102.4942
24.9179 | -102.5455
24.9215 | -102.5954
24.9252 | -102.6458
24.9288 | -102.6934
24.9324 | -102.7383
24.9361 | -102.7843
24.9397 | -102.8294
24.9433 | -102.8897
24.9469 | -102.9517
24.9506 | -103.0011
24.9542 | -103.0334
24.9578 | -103.055
24.9615 | -103.067
24.9651 | -103.0777
24.9687 | -103.0888
24.9723 | -103.0995
24.976 | -103.1083
24.9796 | -103.1152
24.9832 | -103.1201
24.9868 | -103.1247
24.9905 | -103.1277
24.9941 | -103.1305
24.9977 | -103.1317
25.0014 | -103.1317
25.005 | -103.1353
25.0086 | -103.1394
25.0122 | -103.1427
25.0159 | -103.1443
25.0195 | -103.1467
25.0231 | -103.1508
25.0268 | -103.1528
25.0304 | -103.1544
25.034 | -103.1548
25.0376 | -103.1551
25.0413 | -103.1558
25.0449 | -103.1569
25.0485 | -103.1576
25.0522 | -103.1583
25.0558 | -103.1564
25.0594 | -103.1524
25.063 | -103.1489
25.0667 | -103.1473
25.0703 | -103.1455
25.0739 | -103.1404
25.0776 | -103.137
25.0812 | -103.1341
25.0848 | -103.1325
25.0884 | -103.1301
25.0921 | -103.1268
25.0957 | -103.124
25.0993 | -103.1209
25.1029 | -103.1155
25.1066 | -103.1121
25.1102 | -103.1088
25.1138 | -103.1064
25.1175 | -103.103
25.1211 | -103.1007
25.1247 | -103.0978
25.1283 | -103.096
25.132 | -103.0934
25.1356 | -103.0914
25.1392 | -103.0902
25.1429 | -103.0905
25.1465 | -103.0898
25.1501 | -103.0891
25.1537 | -103.0888
25.1574 | -103.0869
25.161 | -103.0859
25.1646 | -103.085
25.1683 | -103.0843
25.1719 | -103.085
25.1755 | -103.0848
25.1791 | -103.084
25.1828 | -103.0845
25.1864 | -103.0847
25.19 | -103.0857
25.1937 | -103.0869
25.1973 | -103.0884
25.2009 | -103.0896
25.2045 | -103.0912
25.2082 | -103.0938
25.2118 | -103.0962
25.2154 | -103.0985
25.219 | -103.102
25.2227 | -103.1042
25.2263 | -103.108
25.2299 | -103.1119
25.2336 | -103.1162
25.2372 | -103.1217
25.2408 | -103.1281
25.2444 | -103.1343
25.2481 | -103.1396
25.2517 | -103.1443
25.2553 | -103.1495
25.259 | -103.1558
25.2626 | -103.162
25.2662 | -103.1698
25.2698 | -103.1784
25.2735 | -103.187
25.2771 | -103.1936
25.2807 | -103.1994
25.2844 | -103.2065
25.288 | -103.2142
25.2916 | -103.2198
25.2952 | -103.2265
25.2989 | -103.2349
25.3025 | -103.2423
25.3061 | -103.2503
25.3098 | -103.254
25.3134 | -103.2578
25.317 | -103.2626
25.3206 | -103.2685
25.3243 | -103.2758
25.3279 | -103.2837
25.3315 | -103.2862
25.3351 | -103.2881
25.3388 | -103.2931
25.3424 | -103.2975
25.346 | -103.3016
25.3497 | -103.3052
25.3533 | -103.3099
25.3569 | -103.3166
25.3605 | -103.3205
25.3642 | -103.3231
25.3678 | -103.3244
25.3714 | -103.3246
25.3751 | -103.3252
25.3787 | -103.3254
25.3823 | -103.3273
25.3859 | -103.33
25.3896 | -103.3323
25.3932 | -103.3314
25.3968 | -103.3333
25.4005 | -103.3344
25.4041 | -103.3347
25.4077 | -103.3351
25.4113 | -103.3348
25.415 | -103.3363
25.4186 | -103.3376
25.4222 | -103.3355
25.4259 | -103.3336
25.4295 | -103.3322
25.4331 | -103.3284
25.4367 | -103.3252
25.4404 | -103.3235
25.444 | -103.3218
25.4476 | -103.3208
25.4512 | -103.3183
25.4549 | -103.3149
25.4585 | -103.3129
25.4621 | -103.3091
25.4658 | -103.3059
25.4694 | -103.3036
25.473 | -103.2996
25.4766 | -103.297
25.4803 | -103.2934
25.4839 | -103.2902
25.4875 | -103.2868
25.4912 | -103.2847
25.4948 | -103.2826
25.4984 | -103.2803
25.502 | -103.2762
25.5057 | -103.2735
25.5093 | -103.2711
25.5129 | -103.2695
25.5166 | -103.2707
25.5202 | -103.2716
25.5238 | -103.2736
25.5274 | -103.2747
25.5311 | -103.2754
25.5347 | -103.2762
25.5383 | -103.2771
25.5419 | -103.278
25.5456 | -103.2785
25.5492 | -103.2804
25.5528 | -103.2834
25.5565 | -103.2844
25.5601 | -103.2856
25.5637 | -103.2891
25.5673 | -103.2913
25.571 | -103.2935
25.5746 | -103.2961
25.5782 | -103.2977
25.5819 | -103.3004
25.5855 | -103.3029
25.5891 | -103.3047
25.5927 | -103.3074
25.5964 | -103.3092
25.6 | -103.3105
25.6036 | -103.3122
25.6073 | -103.3139
25.6109 | -103.3171
25.6145 | -103.3203
25.6181 | -103.3219
25.6218 | -103.3236
25.6254 | -103.3243
25.629 | -103.3248
25.6327 | -103.3257
25.6363 | -103.3278
25.6399 | -103.3272
25.6435 | -103.3274
25.6472 | -103.3236
25.6508 | -103.3239
25.6544 | -103.3255
25.658 | -103.3266
25.6617 | -103.3272
25.6653 | -103.3265
25.6689 | -103.3263
25.6726 | -103.3283
25.6762 | -103.3295
25.6798 | -103.3307
25.6834 | -103.3304
25.6871 | -103.3293
25.6907 | -103.3292
25.6943 | -103.3311
25.698 | -103.3308
25.7016 | -103.3297
25.7052 | -103.3296
25.7088 | -103.3301
25.7125 | -103.3298
25.7161 | -103.3298
25.7197 | -103.3294
25.7234 | -103.3283
25.727 | -103.3289
25.7306 | -103.3294
25.7342 | -103.3322
25.7379 | -103.3343
25.7415 | -103.335
25.7451 | -103.3362
25.7488 | -103.3379
25.7524 | -103.3391
25.756 | -103.3431
25.7596 | -103.3458
25.7633 | -103.3483
25.7669 | -103.3511
25.7705 | -103.3527
25.7741 | -103.3575
25.7778 | -103.3611
25.7814 | -103.3669
25.785 | -103.3737
25.7887 | -103.3805
25.7923 | -103.3875
25.7959 | -103.3931
25.7995 | -103.3981
25.8032 | -103.4036
25.8068 | -103.4094
25.8104 | -103.4146
25.8141 | -103.4229
25.8177 | -103.4326
25.8213 | -103.4416
25.8249 | -103.4492
25.8286 | -103.4588
25.8322 | -103.4686
25.8358 | -103.4799
25.8395 | -103.4907
25.8431 | -103.5013
25.8467 | -103.5129
25.8503 | -103.5242
25.854 | -103.5354
25.8576 | -103.5462
25.8612 | -103.5561
25.8649 | -103.5667
25.8685 | -103.5778
25.8721 | -103.5883
25.8757 | -103.5994
25.8794 | -103.6091
25.883 | -103.6186
25.8866 | -103.6263
25.8902 | -103.6359
25.8939 | -103.6463
25.8975 | -103.6556
25.9011 | -103.6644
25.9048 | -103.6723
25.9084 | -103.6804
25.912 | -103.6869
25.9156 | -103.6919
25.9193 | -103.6988
25.9229 | -103.7057
25.9265 | -103.7109
25.9302 | -103.7161
25.9338 | -103.7206
25.9374 | -103.7257
25.941 | -103.7289
25.9447 | -103.7312
25.9483 | -103.7331
25.9519 | -103.7354
25.9556 | -103.7365
25.9592 | -103.7363
25.9628 | -103.7345
25.9664 | -103.7345
25.9701 | -103.7351
25.9737 | -103.7328
25.9773 | -103.7295
25.981 | -103.7257
25.9846 | -103.7206
25.9882 | -103.7156
25.9918 | -103.7118
25.9955 | -103.7069
25.9991 | -103.7014
26.0027 | -103.6978
26.0063 | -103.6926
26.01 | -103.6882
26.0136 | -103.6821
26.0172 | -103.6751
26.0209 | -103.6678
26.0245 | -103.6591
26.0281 | -103.6498
26.0317 | -103.6427
26.0354 | -103.6346
26.039 | -103.626
26.0426 | -103.6174
26.0463 | -103.6085
26.0499 | -103.6018
26.0535 | -103.5948
26.0571 | -103.5865
26.0608 | -103.5807
26.0644 | -103.5761
26.068 | -103.5703
26.0717 | -103.5641
26.0753 | -103.557
26.0789 | -103.5508
26.0825 | -103.5458
26.0862 | -103.5392
26.0898 | -103.5331
26.0934 | -103.5278
26.0971 | -103.525
26.1007 | -103.5192
26.1043 | -103.5155
26.1079 | -103.5125
26.1116 | -103.5087
26.1152 | -103.5049
26.1188 | -103.5031
26.1224 | -103.5011
26.1261 | -103.4998
26.1297 | -103.4995
26.1333 | -103.4996
26.137 | -103.4994
26.1406 | -103.4973
26.1442 | -103.4962
26.1478 | -103.4925
26.1515 | -103.4893
26.1551 | -103.4893
26.1587 | -103.4898
26.1624 | -103.4916
26.166 | -103.493
26.1696 | -103.4923
26.1732 | -103.4906
26.1769 | -103.4888
26.1805 | -103.4885
26.1841 | -103.4873
26.1878 | -103.4865
26.1914 | -103.4868
26.195 | -103.4881
26.1986 | -103.4876
26.2023 | -103.4859
26.2059 | -103.4858
26.2095 | -103.4864
26.2132 | -103.4864
26.2168 | -103.4855
26.2204 | -103.4844
26.224 | -103.4838
26.2277 | -103.4845
26.2313 | -103.486
26.2349 | -103.4865
26.2385 | -103.4868
26.2422 | -103.487
26.2458 | -103.4857
26.2494 | -103.4854
26.2531 | -103.4851
26.2567 | -103.4853
26.2603 | -103.4845
26.2639 | -103.4822
26.2676 | -103.4797
26.2712 | -103.4782
26.2748 | -103.4778
26.2785 | -103.4775
26.2821 | -103.4771
26.2857 | -103.4755
26.2893 | -103.4731
26.293 | -103.4708
26.2966 | -103.4689
26.3002 | -103.4677
26.3039 | -103.4653
26.3075 | -103.4619
26.3111 | -103.4597
26.3147 | -103.458
26.3184 | -103.4564
26.322 | -103.4545
26.3256 | -103.4523
26.3293 | -103.4514
26.3329 | -103.4504
26.3365 | -103.4495
26.3401 | -103.448
26.3438 | -103.4478
26.3474 | -103.4476
26.351 | -103.4477
26.3546 | -103.4492
26.3583 | -103.4501
26.3619 | -103.4499
26.3655 | -103.4499
26.3692 | -103.4504
26.3728 | -103.452
26.3764 | -103.4534
26.38 | -103.455
26.3837 | -103.4568
26.3873 | -103.458
26.3909 | -103.4592
26.3946 | -103.463
26.3982 | -103.465
26.4018 | -103.4664
26.4054 | -103.4699
26.4091 | -103.4723
26.4127 | -103.4734
26.4163 | -103.4764
26.42 | -103.4763
26.4236 | -103.477
26.4272 | -103.4791
26.4308 | -103.4807
26.4345 | -103.4809
26.4381 | -103.4784
26.4417 | -103.4783
26.4454 | -103.4806
26.449 | -103.4823
26.4526 | -103.4813
26.4562 | -103.482
26.4599 | -103.4834
26.4635 | -103.483
26.4671 | -103.4808
26.4707 | -103.4791
26.4744 | -103.4753
26.478 | -103.4718
26.4816 | -103.4709
26.4853 | -103.4705
26.4889 | -103.468
26.4925 | -103.4648
26.4961 | -103.4606
26.4998 | -103.4555
26.5034 | -103.45
26.507 | -103.4464
26.5107 | -103.4411
26.5143 | -103.4343
26.5179 | -103.4297
26.5215 | -103.4246
26.5252 | -103.4186
26.5288 | -103.4123
26.5324 | -103.4054
26.5361 | -103.3981
26.5397 | -103.3907
26.5433 | -103.3835
26.5469 | -103.3766
26.5506 | -103.3695
26.5542 | -103.3605
26.5578 | -103.3513
26.5615 | -103.3429
26.5651 | -103.3327
26.5687 | -103.3227
26.5723 | -103.3155
26.576 | -103.3089
26.5796 | -103.3002
26.5832 | -103.2924
26.5868 | -103.2861
26.5905 | -103.2796
26.5941 | -103.2734
26.5977 | -103.2671
26.6014 | -103.2605
26.605 | -103.2553
26.6086 | -103.2507
26.6122 | -103.2467
26.6159 | -103.2426
26.6195 | -103.2389
26.6231 | -103.234
26.6268 | -103.229
26.6304 | -103.2256
26.634 | -103.2228
26.6376 | -103.221
26.6413 | -103.2214
26.6449 | -103.2213
26.6485 | -103.2199
26.6522 | -103.2196
26.6558 | -103.2198
26.6594 | -103.2205
26.663 | -103.2228
26.6667 | -103.226
26.6703 | -103.2279
26.6739 | -103.23
26.6776 | -103.2343
26.6812 | -103.2369
26.6848 | -103.2389
26.6884 | -103.2433
26.6921 | -103.246
26.6957 | -103.2502
26.6993 | -103.2552
26.7029 | -103.2608
26.7066 | -103.2668
26.7102 | -103.2724
26.7138 | -103.2778
26.7175 | -103.2828
26.7211 | -103.2861
26.7247 | -103.2901
26.7283 | -103.2945
26.732 | -103.2977
26.7356 | -103.3023
26.7392 | -103.3072
26.7429 | -103.312
26.7465 | -103.3158
26.7501 | -103.3182
26.7537 | -103.3213
26.7574 | -103.3236
26.761 | -103.3258
26.7646 | -103.3283
26.7683 | -103.3319
26.7719 | -103.3338
26.7755 | -103.3355
26.7791 | -103.3377
26.7828 | -103.3383
26.7864 | -103.3396
26.79 | -103.34
26.7937 | -103.3387
26.7973 | -103.3394
26.8009 | -103.3387
26.8045 | -103.3354
26.8082 | -103.3303
26.8118 | -103.3289
26.8154 | -103.328
26.819 | -103.3215
26.8227 | -103.3142
26.8263 | -103.3088
26.8299 | -103.3028
26.8336 | -103.2943
26.8372 | -103.285
26.8408 | -103.276
26.8444 | -103.2667
26.8481 | -103.2568
26.8517 | -103.2476
26.8553 | -103.238
26.859 | -103.2232
26.8626 | -103.2083
26.8662 | -103.1946
26.8698 | -103.1827
26.8735 | -103.1709
26.8771 | -103.1571
26.8807 | -103.1439
26.8844 | -103.1298
26.888 | -103.1165
26.8916 | -103.1033
26.8952 | -103.0876
26.8989 | -103.0705
26.9025 | -103.0518
26.9061 | -103.0328
26.9098 | -103.0171
26.9134 | -103.0023
26.917 | -102.987
26.9206 | -102.9712
26.9243 | -102.9538
26.9279 | -102.9374
26.9315 | -102.9207
26.9351 | -102.9014
26.9388 | -102.882
26.9424 | -102.8631
26.946 | -102.8451
26.9497 | -102.8271
26.9533 | -102.8085
26.9569 | -102.7907
26.9605 | -102.7725
26.9642 | -102.7534
26.9678 | -102.7358
26.9714 | -102.7183
26.9751 | -102.7024
26.9787 | -102.6869
26.9823 | -102.671
26.9859 | -102.6559
26.9896 | -102.6415
26.9932 | -102.6279
26.9968 | -102.6144
27.0005 | -102.6016
27.0041 | -102.5907
27.0077 | -102.5791
27.0113 | -102.5677
27.015 | -102.5569
27.0186 | -102.5467
27.0222 | -102.5359
27.0258 | -102.525
27.0295 | -102.5162
27.0331 | -102.5058
27.0367 | -102.4979
27.0404 | -102.4902
27.044 | -102.4826
27.0476 | -102.4763
27.0512 | -102.4724
27.0549 | -102.4695
27.0585 | -102.4684
27.0621 | -102.4686
27.0658 | -102.4676
27.0694 | -102.4702
27.073 | -102.4722
27.0766 | -102.4732
27.0803 | -102.4749
27.0839 | -102.477
27.0875 | -102.4791
27.0912 | -102.4841
27.0948 | -102.4896
27.0984 | -102.495
27.102 | -102.5041
27.1057 | -102.5135
27.1093 | -102.5221
27.1129 | -102.5336
27.1166 | -102.545
27.1202 | -102.5562
27.1238 | -102.5682
27.1274 | -102.5817
27.1311 | -102.5956
27.1347 | -102.6097
27.1383 | -102.6246
27.142 | -102.6399
27.1456 | -102.6565
27.1492 | -102.6711
27.1528 | -102.6834
27.1565 | -102.6985
27.1601 | -102.7135
27.1637 | -102.7301
27.1673 | -102.747
27.171 | -102.7647
27.1746 | -102.7818
27.1782 | -102.7985
27.1819 | -102.814
27.1855 | -102.8311
27.1891 | -102.8488
27.1927 | -102.8641
27.1964 | -102.8806
27.2 | -102.8977
27.2036 | -102.9125
27.2073 | -102.9253
27.2109 | -102.9395
27.2145 | -102.9509
27.2181 | -102.9665
27.2218 | -102.9799
27.2254 | -102.9913
27.229 | -103.0031
27.2327 | -103.0155
27.2363 | -103.0267
27.2399 | -103.0372
27.2435 | -103.0484
27.2472 | -103.0589
27.2508 | -103.0684
27.2544 | -103.076
27.2581 | -103.0847
27.2617 | -103.0941
27.2653 | -103.1024
27.2689 | -103.11
27.2726 | -103.1182
27.2762 | -103.1263
27.2798 | -103.1335
27.2834 | -103.1398
27.2871 | -103.1469
27.2907 | -103.1535
27.2943 | -103.158
27.298 | -103.161
27.3016 | -103.1637
27.3052 | -103.1676
27.3088 | -103.1726
27.3125 | -103.1758
27.3161 | -103.1793
27.3197 | -103.1853
27.3234 | -103.1902
27.327 | -103.1945
27.3306 | -103.1974
27.3342 | -103.2
27.3379 | -103.2024
27.3415 | -103.204
27.3451 | -103.2055
27.3488 | -103.2087
27.3524 | -103.2119
27.356 | -103.2142
27.3596 | -103.2156
27.3633 | -103.2167
27.3669 | -103.2169
27.3705 | -103.2178
27.3742 | -103.2183
27.3778 | -103.2202
27.3814 | -103.221
27.385 | -103.2237
27.3887 | -103.2246
27.3923 | -103.2244
27.3959 | -103.2259
27.3995 | -103.2286
27.4032 | -103.2309
27.4068 | -103.2347
27.4104 | -103.2411
27.4141 | -103.245
27.4177 | -103.2506
27.4213 | -103.2544
27.4249 | -103.2578
27.4286 | -103.26
27.4322 | -103.2616
27.4358 | -103.2627
27.4395 | -103.2641
27.4431 | -103.2661
27.4467 | -103.2685
27.4503 | -103.2711
27.454 | -103.2736
27.4576 | -103.2754
27.4612 | -103.2777
27.4649 | -103.2798
27.4685 | -103.2826
27.4721 | -103.2856
27.4757 | -103.2874
27.4794 | -103.2898
27.483 | -103.2933
27.4866 | -103.2957
27.4902 | -103.2993
27.4939 | -103.303
27.4975 | -103.3029
27.5011 | -103.3037
27.5048 | -103.3069
27.5084 | -103.3102
27.512 | -103.3127
27.5156 | -103.3168
27.5193 | -103.3212
27.5229 | -103.3254
27.5265 | -103.3305
27.5302 | -103.3332
27.5338 | -103.3375
27.5374 | -103.3461
27.541 | -103.351
27.5447 | -103.3558
27.5483 | -103.3605
27.5519 | -103.3657
27.5556 | -103.3705
27.5592 | -103.3758
27.5628 | -103.3803
27.5664 | -103.3871
27.5701 | -103.3947
27.5737 | -103.4011
27.5773 | -103.4063
27.581 | -103.4116
27.5846 | -103.4178
27.5882 | -103.4238
27.5918 | -103.4316
27.5955 | -103.4388
27.5991 | -103.4451
27.6027 | -103.4514
27.6063 | -103.4578
27.61 | -103.4656
27.6136 | -103.4743
27.6172 | -103.4818
27.6209 | -103.4877
27.6245 | -103.493
27.6281 | -103.4988
27.6317 | -103.5064
27.6354 | -103.5128
27.639 | -103.5183
27.6426 | -103.5245
27.6463 | -103.5297
27.6499 | -103.5356
27.6535 | -103.5408
27.6571 | -103.5459
27.6608 | -103.5505
27.6644 | -103.5546
27.668 | -103.5592
27.6717 | -103.5648
27.6753 | -103.5702
27.6789 | -103.5744
27.6825 | -103.5772
27.6862 | -103.5823
27.6898 | -103.5873
27.6934 | -103.5903
27.6971 | -103.5931
27.7007 | -103.5956
27.7043 | -103.5973
27.7079 | -103.599
27.7116 | -103.6001
27.7152 | -103.6024
27.7188 | -103.6051
27.7224 | -103.6056
27.7261 | -103.6074
27.7297 | -103.6105
27.7333 | -103.6124
27.737 | -103.6139
27.7406 | -103.6142
27.7442 | -103.6142
27.7478 | -103.6154
27.7515 | -103.617
27.7551 | -103.6175
27.7587 | -103.6179
27.7624 | -103.6171
27.766 | -103.6155
27.7696 | -103.6137
27.7732 | -103.6141
27.7769 | -103.6125
27.7805 | -103.611
27.7841 | -103.6117
27.7878 | -103.6106
27.7914 | -103.6093
27.795 | -103.6074
27.7986 | -103.6057
27.8023 | -103.6027
27.8059 | -103.601
27.8095 | -103.599
27.8132 | -103.5981
27.8168 | -103.5977
27.8204 | -103.598
27.824 | -103.5985
27.8277 | -103.5976
27.8313 | -103.5967
27.8349 | -103.5949
27.8385 | -103.5942
27.8422 | -103.5929
27.8458 | -103.5915
27.8494 | -103.5904
27.8531 | -103.5901
27.8567 | -103.5883
27.8603 | -103.5865
27.8639 | -103.5853
27.8676 | -103.582
27.8712 | -103.5785
27.8748 | -103.5757
27.8785 | -103.5726
27.8821 | -103.567
27.8857 | -103.5603
27.8893 | -103.5583
27.893 | -103.555
27.8966 | -103.5505
27.9002 | -103.5463
27.9039 | -103.5423
27.9075 | -103.5371
27.9111 | -103.5304
27.9147 | -103.5239
27.9184 | -103.5182
27.922 | -103.5129
27.9256 | -103.5069
27.9293 | -103.498
27.9329 | -103.4913
27.9365 | -103.4838
27.9401 | -103.4759
27.9438 | -103.4688
27.9474 | -103.4605
27.951 | -103.4524
27.9546 | -103.4452
27.9583 | -103.436
27.9619 | -103.4254
27.9655 | -103.4168
27.9692 | -103.4093
27.9728 | -103.4011
27.9764 | -103.3921
27.98 | -103.3846
27.9837 | -103.3751
27.9873 | -103.3652
27.9909 | -103.3569
27.9946 | -103.3477
27.9982 | -103.3389
28.0018 | -103.3312