## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 226 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 229
Hood | 02-92B/P | 168
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1013.9 | hPa |  | PRC PRO 01
Humidity | 41.7 | %/Hr
Internal temperature of calibration bench | 22.82 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1020 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 23.6 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.95 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.02 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3108 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 15/05/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 16/05/2018
Engineer | Anthony HUE | Approval |  | 18/05/2018
CEO | Robert GLEMOT | Approval |  | 18/05/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3797 | 57.9719
0.012209 | -35.9947 | 38.9415
0.018313 | -35.0069 | 28.8181
0.024417 | -34.5645 | 22.2493
0.030521 | -34.3893 | 18.23
0.036626 | -34.2405 | 15.1499
0.04273 | -34.1556 | 13.0171
0.048834 | -34.0964 | 11.3169
0.054938 | -34.0552 | 10.005
0.061043 | -34.0114 | 8.945
0.067147 | -33.9947 | 8.0876
0.073251 | -33.9814 | 7.3448
0.079355 | -33.9734 | 6.8037
0.08546 | -33.9713 | 6.3903
0.091564 | -33.9667 | 6.0009
0.097668 | -33.9453 | 5.6189
0.10377 | -33.9136 | 5.1999
0.10988 | -33.8948 | 4.8027
0.11598 | -33.892 | 4.4515
0.12209 | -33.9066 | 4.1979
0.12819 | -33.9304 | 4.0081
0.13429 | -33.9521 | 3.8266
0.1404 | -33.9655 | 3.6233
0.1465 | -33.9787 | 3.4225
0.15261 | -33.9877 | 3.2213
0.15871 | -33.9896 | 3.027
0.16482 | -33.9823 | 2.8323
0.17092 | -33.9599 | 2.6975
0.17702 | -33.9428 | 2.5548
0.18313 | -33.9262 | 2.4385
0.18923 | -33.9138 | 2.3304
0.19534 | -33.9077 | 2.2444
0.20144 | -33.9075 | 2.1652
0.20754 | -33.9127 | 2.1163
0.21365 | -33.9159 | 2.0611
0.21975 | -33.9206 | 2.0057
0.22586 | -33.9278 | 1.9447
0.23196 | -33.9344 | 1.8669
0.23807 | -33.9398 | 1.7786
0.24417 | -33.9387 | 1.6763
0.25027 | -33.9388 | 1.5853
0.25638 | -33.9369 | 1.4889
0.26248 | -33.9337 | 1.4034
0.26859 | -33.9309 | 1.3221
0.27469 | -33.9302 | 1.2447
0.2808 | -33.9357 | 1.1957
0.2869 | -33.9399 | 1.1396
0.293 | -33.945 | 1.0928
0.29911 | -33.951 | 1.0608
0.30521 | -33.9567 | 1.0314
0.31132 | -33.9606 | 1.0112
0.31742 | -33.9644 | 0.98518
0.32353 | -33.9679 | 0.95799
0.32963 | -33.9705 | 0.93022
0.33573 | -33.9719 | 0.89793
0.34184 | -33.9735 | 0.86466
0.34794 | -33.9751 | 0.8173
0.35405 | -33.9768 | 0.77444
0.36015 | -33.9784 | 0.73006
0.36626 | -33.9792 | 0.68482
0.37236 | -33.98 | 0.6355
0.37846 | -33.9805 | 0.58579
0.38457 | -33.979 | 0.54743
0.39067 | -33.9789 | 0.50816
0.39678 | -33.9768 | 0.46797
0.40288 | -33.9743 | 0.43319
0.40899 | -33.9709 | 0.39856
0.41509 | -33.9672 | 0.36302
0.42119 | -33.9626 | 0.32716
0.4273 | -33.9591 | 0.2917
0.4334 | -33.9561 | 0.25574
0.43951 | -33.9531 | 0.21861
0.44561 | -33.9511 | 0.18186
0.45172 | -33.9516 | 0.14734
0.45782 | -33.951 | 0.11286
0.46392 | -33.9515 | 0.081435
0.47003 | -33.9526 | 0.050627
0.47613 | -33.9543 | 0.02463
0.48224 | -33.9564 | 0.0013003
0.48834 | -33.9601 | -0.011613
0.49445 | -33.9627 | -0.027884
0.50055 | -33.9662 | -0.044457
0.50665 | -33.969 | -0.054287
0.51276 | -33.9714 | -0.065563
0.51886 | -33.974 | -0.073911
0.52497 | -33.9745 | -0.078944
0.53107 | -33.976 | -0.083484
0.53717 | -33.9774 | -0.090036
0.54328 | -33.9783 | -0.094803
0.54938 | -33.9787 | -0.10266
0.55549 | -33.9786 | -0.11259
0.56159 | -33.9779 | -0.13223
0.5677 | -33.9775 | -0.14895
0.5738 | -33.9764 | -0.16675
0.5799 | -33.9761 | -0.18443
0.58601 | -33.9758 | -0.21073
0.59211 | -33.9757 | -0.24596
0.59822 | -33.9754 | -0.27267
0.60432 | -33.9754 | -0.30189
0.61043 | -33.9755 | -0.33425
0.61653 | -33.9759 | -0.36472
0.62263 | -33.9759 | -0.39733
0.62874 | -33.977 | -0.42552
0.63484 | -33.9776 | -0.45607
0.64095 | -33.9787 | -0.48554
0.64705 | -33.9798 | -0.51422
0.65316 | -33.9809 | -0.53989
0.65926 | -33.9822 | -0.56551
0.66536 | -33.9829 | -0.58489
0.67147 | -33.9838 | -0.6063
0.67757 | -33.985 | -0.62814
0.68368 | -33.9858 | -0.64781
0.68978 | -33.9869 | -0.66692
0.69589 | -33.9876 | -0.68451
0.70199 | -33.988 | -0.70167
0.70809 | -33.9886 | -0.7207
0.7142 | -33.9891 | -0.73743
0.7203 | -33.9895 | -0.75553
0.72641 | -33.9896 | -0.7717
0.73251 | -33.9903 | -0.78997
0.73862 | -33.9904 | -0.80952
0.74472 | -33.9913 | -0.82598
0.75082 | -33.9914 | -0.84215
0.75693 | -33.9915 | -0.86095
0.76303 | -33.9913 | -0.87839
0.76914 | -33.9904 | -0.90127
0.77524 | -33.9901 | -0.92318
0.78135 | -33.9896 | -0.94296
0.78745 | -33.9893 | -0.96436
0.79355 | -33.9889 | -0.98557
0.79966 | -33.9886 | -1.0082
0.80576 | -33.9879 | -1.0349
0.81187 | -33.987 | -1.0593
0.81797 | -33.9862 | -1.0851
0.82408 | -33.9855 | -1.108
0.83018 | -33.9847 | -1.1348
0.83628 | -33.9841 | -1.1628
0.84239 | -33.9831 | -1.1864
0.84849 | -33.9821 | -1.209
0.8546 | -33.9814 | -1.2329
0.8607 | -33.9806 | -1.2574
0.86681 | -33.9799 | -1.2827
0.87291 | -33.9788 | -1.3053
0.87901 | -33.9781 | -1.3252
0.88512 | -33.9775 | -1.3457
0.89122 | -33.9769 | -1.3663
0.89733 | -33.9766 | -1.3852
0.90343 | -33.9761 | -1.4043
0.90953 | -33.9764 | -1.4195
0.91564 | -33.976 | -1.4348
0.92174 | -33.9757 | -1.4489
0.92785 | -33.9754 | -1.4627
0.93395 | -33.9755 | -1.4752
0.94006 | -33.9756 | -1.4905
0.94616 | -33.976 | -1.4973
0.95226 | -33.9762 | -1.5105
0.95837 | -33.9765 | -1.5228
0.96447 | -33.977 | -1.5344
0.97058 | -33.9774 | -1.5453
0.97668 | -33.9776 | -1.5542
0.98279 | -33.9782 | -1.559
0.98889 | -33.9788 | -1.5669
0.99499 | -33.9792 | -1.5761
1.0011 | -33.9798 | -1.584
1.0072 | -33.9802 | -1.5898
1.0133 | -33.9807 | -1.5966
1.0194 | -33.9813 | -1.6026
1.0255 | -33.9816 | -1.6097
1.0316 | -33.9817 | -1.6155
1.0377 | -33.9818 | -1.6214
1.0438 | -33.9818 | -1.6268
1.0499 | -33.9811 | -1.6316
1.056 | -33.9808 | -1.6358
1.0621 | -33.9804 | -1.6409
1.0682 | -33.9806 | -1.6468
1.0743 | -33.9802 | -1.6529
1.0805 | -33.9803 | -1.6589
1.0866 | -33.9796 | -1.668
1.0927 | -33.9792 | -1.6745
1.0988 | -33.9792 | -1.6838
1.1049 | -33.9786 | -1.6906
1.111 | -33.978 | -1.6998
1.1171 | -33.9774 | -1.708
1.1232 | -33.9766 | -1.7203
1.1293 | -33.976 | -1.7315
1.1354 | -33.9755 | -1.7422
1.1415 | -33.9748 | -1.7541
1.1476 | -33.9742 | -1.766
1.1537 | -33.9734 | -1.7779
1.1598 | -33.9734 | -1.7945
1.1659 | -33.9732 | -1.8062
1.172 | -33.9729 | -1.8187
1.1781 | -33.9726 | -1.8319
1.1842 | -33.9723 | -1.8464
1.1903 | -33.9729 | -1.8633
1.1964 | -33.9731 | -1.8774
1.2025 | -33.9731 | -1.8917
1.2086 | -33.9734 | -1.9078
1.2147 | -33.9739 | -1.9231
1.2209 | -33.9744 | -1.938
1.227 | -33.9759 | -1.9548
1.2331 | -33.9767 | -1.9697
1.2392 | -33.9775 | -1.9865
1.2453 | -33.9786 | -2.0001
1.2514 | -33.9797 | -2.0148
1.2575 | -33.981 | -2.0304
1.2636 | -33.9828 | -2.0452
1.2697 | -33.9842 | -2.06
1.2758 | -33.9853 | -2.0756
1.2819 | -33.9867 | -2.0915
1.288 | -33.988 | -2.107
1.2941 | -33.9895 | -2.1214
1.3002 | -33.9911 | -2.1336
1.3063 | -33.9925 | -2.1468
1.3124 | -33.9939 | -2.16
1.3185 | -33.9955 | -2.1738
1.3246 | -33.9969 | -2.1877
1.3307 | -33.9985 | -2.2007
1.3368 | -33.9998 | -2.2114
1.3429 | -34.0011 | -2.2234
1.349 | -34.0023 | -2.2348
1.3551 | -34.0036 | -2.2485
1.3613 | -34.0049 | -2.2601
1.3674 | -34.0058 | -2.2705
1.3735 | -34.0069 | -2.2809
1.3796 | -34.0079 | -2.2913
1.3857 | -34.0088 | -2.301
1.3918 | -34.0096 | -2.3126
1.3979 | -34.0106 | -2.3238
1.404 | -34.011 | -2.3345
1.4101 | -34.0118 | -2.3451
1.4162 | -34.0124 | -2.3552
1.4223 | -34.013 | -2.3662
1.4284 | -34.0135 | -2.3764
1.4345 | -34.0141 | -2.3873
1.4406 | -34.0145 | -2.397
1.4467 | -34.015 | -2.4069
1.4528 | -34.0154 | -2.4167
1.4589 | -34.016 | -2.4274
1.465 | -34.0162 | -2.4362
1.4711 | -34.0169 | -2.4457
1.4772 | -34.0171 | -2.4565
1.4833 | -34.0175 | -2.4678
1.4894 | -34.018 | -2.479
1.4955 | -34.0184 | -2.4905
1.5016 | -34.0187 | -2.5019
1.5078 | -34.019 | -2.5147
1.5139 | -34.0195 | -2.5262
1.52 | -34.0198 | -2.5388
1.5261 | -34.0201 | -2.5497
1.5322 | -34.0202 | -2.5605
1.5383 | -34.0203 | -2.572
1.5444 | -34.0207 | -2.5844
1.5505 | -34.0209 | -2.5963
1.5566 | -34.0211 | -2.608
1.5627 | -34.0215 | -2.6195
1.5688 | -34.0219 | -2.631
1.5749 | -34.022 | -2.6425
1.581 | -34.0223 | -2.6578
1.5871 | -34.0227 | -2.6716
1.5932 | -34.0229 | -2.6835
1.5993 | -34.0231 | -2.6982
1.6054 | -34.0233 | -2.7118
1.6115 | -34.0235 | -2.7267
1.6176 | -34.0238 | -2.7423
1.6237 | -34.024 | -2.755
1.6298 | -34.0242 | -2.7697
1.6359 | -34.0245 | -2.7833
1.642 | -34.0247 | -2.7975
1.6482 | -34.0251 | -2.8134
1.6543 | -34.0255 | -2.8284
1.6604 | -34.0258 | -2.8428
1.6665 | -34.026 | -2.8571
1.6726 | -34.0262 | -2.8726
1.6787 | -34.0266 | -2.8888
1.6848 | -34.0267 | -2.9045
1.6909 | -34.0271 | -2.9198
1.697 | -34.0273 | -2.9356
1.7031 | -34.0274 | -2.9505
1.7092 | -34.0277 | -2.965
1.7153 | -34.0279 | -2.981
1.7214 | -34.0278 | -2.996
1.7275 | -34.0279 | -3.0111
1.7336 | -34.0279 | -3.0263
1.7397 | -34.0279 | -3.041
1.7458 | -34.0282 | -3.0572
1.7519 | -34.0281 | -3.0729
1.758 | -34.0282 | -3.0876
1.7641 | -34.0283 | -3.1021
1.7702 | -34.0283 | -3.1168
1.7763 | -34.0285 | -3.1306
1.7824 | -34.0285 | -3.1462
1.7885 | -34.0285 | -3.1596
1.7947 | -34.0287 | -3.1741
1.8008 | -34.0288 | -3.1874
1.8069 | -34.0288 | -3.201
1.813 | -34.0289 | -3.2147
1.8191 | -34.0288 | -3.2281
1.8252 | -34.0287 | -3.241
1.8313 | -34.0288 | -3.2549
1.8374 | -34.0288 | -3.2682
1.8435 | -34.029 | -3.2801
1.8496 | -34.0288 | -3.2933
1.8557 | -34.0287 | -3.3071
1.8618 | -34.0286 | -3.3183
1.8679 | -34.0287 | -3.331
1.874 | -34.0287 | -3.3438
1.8801 | -34.0286 | -3.3561
1.8862 | -34.0287 | -3.3684
1.8923 | -34.0287 | -3.3801
1.8984 | -34.0284 | -3.3903
1.9045 | -34.0284 | -3.4013
1.9106 | -34.0282 | -3.4126
1.9167 | -34.028 | -3.4233
1.9228 | -34.028 | -3.4339
1.9289 | -34.0278 | -3.4425
1.9351 | -34.0278 | -3.4521
1.9412 | -34.0278 | -3.4626
1.9473 | -34.0278 | -3.4741
1.9534 | -34.0277 | -3.4852
1.9595 | -34.0274 | -3.4965
1.9656 | -34.0269 | -3.5064
1.9717 | -34.0266 | -3.5167
1.9778 | -34.0265 | -3.5282
1.9839 | -34.0265 | -3.5396
1.99 | -34.0265 | -3.5498
1.9961 | -34.0263 | -3.5596
2.0022 | -34.0262 | -3.5696
2.0083 | -34.0259 | -3.5797
2.0144 | -34.0258 | -3.5895
2.0205 | -34.0257 | -3.601
2.0266 | -34.0257 | -3.6126
2.0327 | -34.0257 | -3.6236
2.0388 | -34.0257 | -3.6372
2.0449 | -34.0257 | -3.6475
2.051 | -34.0257 | -3.6575
2.0571 | -34.0254 | -3.6688
2.0632 | -34.0253 | -3.6793
2.0693 | -34.0253 | -3.6897
2.0754 | -34.0254 | -3.7015
2.0816 | -34.0254 | -3.7128
2.0877 | -34.0254 | -3.7237
2.0938 | -34.0255 | -3.7353
2.0999 | -34.0254 | -3.7466
2.106 | -34.0256 | -3.7583
2.1121 | -34.0256 | -3.77
2.1182 | -34.0257 | -3.7805
2.1243 | -34.0258 | -3.7916
2.1304 | -34.0259 | -3.802
2.1365 | -34.0261 | -3.8132
2.1426 | -34.0262 | -3.8259
2.1487 | -34.0266 | -3.8376
2.1548 | -34.0268 | -3.8502
2.1609 | -34.0272 | -3.862
2.167 | -34.0274 | -3.8747
2.1731 | -34.0277 | -3.8859
2.1792 | -34.0283 | -3.8974
2.1853 | -34.0287 | -3.9097
2.1914 | -34.0291 | -3.9219
2.1975 | -34.0295 | -3.9338
2.2036 | -34.0299 | -3.9463
2.2097 | -34.0303 | -3.9582
2.2158 | -34.0311 | -3.9689
2.222 | -34.0315 | -3.9798
2.2281 | -34.0318 | -3.9905
2.2342 | -34.0322 | -4.0036
2.2403 | -34.0327 | -4.0142
2.2464 | -34.0334 | -4.0257
2.2525 | -34.0338 | -4.0364
2.2586 | -34.0346 | -4.0468
2.2647 | -34.0353 | -4.0568
2.2708 | -34.0358 | -4.0685
2.2769 | -34.0365 | -4.0805
2.283 | -34.037 | -4.09
2.2891 | -34.0375 | -4.1006
2.2952 | -34.0381 | -4.111
2.3013 | -34.0388 | -4.1215
2.3074 | -34.0394 | -4.1319
2.3135 | -34.0398 | -4.1431
2.3196 | -34.0405 | -4.1526
2.3257 | -34.0408 | -4.1637
2.3318 | -34.0412 | -4.1729
2.3379 | -34.0417 | -4.1848
2.344 | -34.0421 | -4.1955
2.3501 | -34.0425 | -4.2067
2.3562 | -34.0429 | -4.2157
2.3623 | -34.0432 | -4.2251
2.3685 | -34.0435 | -4.2348
2.3746 | -34.0438 | -4.2453
2.3807 | -34.0443 | -4.2549
2.3868 | -34.0445 | -4.2643
2.3929 | -34.0449 | -4.2741
2.399 | -34.0453 | -4.2828
2.4051 | -34.0457 | -4.2928
2.4112 | -34.046 | -4.3023
2.4173 | -34.0463 | -4.3117
2.4234 | -34.0464 | -4.3205
2.4295 | -34.0468 | -4.3293
2.4356 | -34.0469 | -4.339
2.4417 | -34.047 | -4.349
2.4478 | -34.0471 | -4.3589
2.4539 | -34.0475 | -4.3675
2.46 | -34.0472 | -4.3769
2.4661 | -34.0473 | -4.3862
2.4722 | -34.0475 | -4.3957
2.4783 | -34.0477 | -4.4048
2.4844 | -34.0478 | -4.4134
2.4905 | -34.0478 | -4.4228
2.4966 | -34.0476 | -4.432
2.5027 | -34.0476 | -4.4406
2.5089 | -34.0475 | -4.4493
2.515 | -34.0475 | -4.4588
2.5211 | -34.0475 | -4.4674
2.5272 | -34.0473 | -4.4766
2.5333 | -34.0471 | -4.4863
2.5394 | -34.0471 | -4.4955
2.5455 | -34.0469 | -4.5048
2.5516 | -34.0469 | -4.5141
2.5577 | -34.0468 | -4.5238
2.5638 | -34.0467 | -4.5334
2.5699 | -34.0466 | -4.5434
2.576 | -34.0464 | -4.5534
2.5821 | -34.0463 | -4.5632
2.5882 | -34.0461 | -4.5727
2.5943 | -34.0461 | -4.5827
2.6004 | -34.0459 | -4.5928
2.6065 | -34.0457 | -4.6018
2.6126 | -34.0457 | -4.6114
2.6187 | -34.0456 | -4.6206
2.6248 | -34.0455 | -4.6299
2.6309 | -34.0453 | -4.6408
2.637 | -34.0453 | -4.6513
2.6431 | -34.0453 | -4.6613
2.6492 | -34.0452 | -4.671
2.6554 | -34.0452 | -4.6814
2.6615 | -34.045 | -4.6913
2.6676 | -34.0449 | -4.7014
2.6737 | -34.045 | -4.7119
2.6798 | -34.0449 | -4.7217
2.6859 | -34.0449 | -4.7319
2.692 | -34.0448 | -4.7417
2.6981 | -34.0447 | -4.7528
2.7042 | -34.0448 | -4.7647
2.7103 | -34.0447 | -4.7751
2.7164 | -34.0449 | -4.7858
2.7225 | -34.0449 | -4.7968
2.7286 | -34.0449 | -4.8079
2.7347 | -34.0449 | -4.8183
2.7408 | -34.0451 | -4.8282
2.7469 | -34.045 | -4.8394
2.753 | -34.0452 | -4.8503
2.7591 | -34.0452 | -4.8619
2.7652 | -34.0452 | -4.873
2.7713 | -34.0452 | -4.8831
2.7774 | -34.0455 | -4.8944
2.7835 | -34.0457 | -4.9051
2.7896 | -34.0458 | -4.9147
2.7958 | -34.0461 | -4.9251
2.8019 | -34.0465 | -4.9354
2.808 | -34.0467 | -4.9474
2.8141 | -34.047 | -4.9601
2.8202 | -34.0473 | -4.9723
2.8263 | -34.0476 | -4.9832
2.8324 | -34.0477 | -4.9952
2.8385 | -34.0479 | -5.0073
2.8446 | -34.0483 | -5.0191
2.8507 | -34.0485 | -5.0313
2.8568 | -34.0488 | -5.0427
2.8629 | -34.0492 | -5.0544
2.869 | -34.0495 | -5.0655
2.8751 | -34.0498 | -5.0776
2.8812 | -34.0503 | -5.0905
2.8873 | -34.0505 | -5.1019
2.8934 | -34.0508 | -5.1136
2.8995 | -34.051 | -5.1254
2.9056 | -34.0513 | -5.136
2.9117 | -34.0516 | -5.1474
2.9178 | -34.0521 | -5.1602
2.9239 | -34.0524 | -5.1729
2.93 | -34.0527 | -5.1858
2.9361 | -34.0531 | -5.1974
2.9423 | -34.0536 | -5.2096
2.9484 | -34.0539 | -5.2229
2.9545 | -34.0544 | -5.236
2.9606 | -34.0548 | -5.2489
2.9667 | -34.0551 | -5.2605
2.9728 | -34.0554 | -5.2724
2.9789 | -34.0557 | -5.2849
2.985 | -34.0559 | -5.2983
2.9911 | -34.0563 | -5.3105
2.9972 | -34.0566 | -5.3229
3.0033 | -34.057 | -5.3354
3.0094 | -34.0574 | -5.3482
3.0155 | -34.0578 | -5.3608
3.0216 | -34.0581 | -5.3738
3.0277 | -34.0585 | -5.3869
3.0338 | -34.0589 | -5.4
3.0399 | -34.0593 | -5.4132
3.046 | -34.0595 | -5.4257
3.0521 | -34.06 | -5.438
3.0582 | -34.0602 | -5.4508
3.0643 | -34.0607 | -5.4634
3.0704 | -34.0608 | -5.476
3.0765 | -34.0611 | -5.488
3.0827 | -34.0614 | -5.5009
3.0888 | -34.0616 | -5.5134
3.0949 | -34.0618 | -5.5259
3.101 | -34.0621 | -5.5395
3.1071 | -34.0623 | -5.5528
3.1132 | -34.0626 | -5.5661
3.1193 | -34.0629 | -5.5786
3.1254 | -34.0632 | -5.5902
3.1315 | -34.0635 | -5.6037
3.1376 | -34.0637 | -5.6161
3.1437 | -34.0639 | -5.629
3.1498 | -34.0641 | -5.6421
3.1559 | -34.0644 | -5.6549
3.162 | -34.0644 | -5.6679
3.1681 | -34.0644 | -5.6812
3.1742 | -34.0647 | -5.6943
3.1803 | -34.065 | -5.7068
3.1864 | -34.0652 | -5.7187
3.1925 | -34.0655 | -5.7318
3.1986 | -34.0655 | -5.7436
3.2047 | -34.0656 | -5.7566
3.2108 | -34.0658 | -5.7681
3.2169 | -34.0661 | -5.7811
3.223 | -34.0661 | -5.7939
3.2292 | -34.066 | -5.807
3.2353 | -34.066 | -5.8166
3.2414 | -34.0661 | -5.8291
3.2475 | -34.066 | -5.8396
3.2536 | -34.0662 | -5.8529
3.2597 | -34.0664 | -5.864
3.2658 | -34.0666 | -5.8746
3.2719 | -34.0667 | -5.8855
3.278 | -34.0667 | -5.897
3.2841 | -34.0669 | -5.9087
3.2902 | -34.0669 | -5.9207
3.2963 | -34.067 | -5.9319
3.3024 | -34.067 | -5.9419
3.3085 | -34.0672 | -5.9538
3.3146 | -34.0671 | -5.9644
3.3207 | -34.0673 | -5.976
3.3268 | -34.0674 | -5.9877
3.3329 | -34.0675 | -5.9976
3.339 | -34.0677 | -6.0072
3.3451 | -34.0677 | -6.0184
3.3512 | -34.0679 | -6.0279
3.3573 | -34.068 | -6.0389
3.3634 | -34.0681 | -6.0501
3.3696 | -34.0681 | -6.0616
3.3757 | -34.0682 | -6.0715
3.3818 | -34.0684 | -6.0831
3.3879 | -34.0684 | -6.0932
3.394 | -34.0684 | -6.1032
3.4001 | -34.0685 | -6.1141
3.4062 | -34.0686 | -6.1242
3.4123 | -34.0687 | -6.1342
3.4184 | -34.0687 | -6.1448
3.4245 | -34.0688 | -6.1553
3.4306 | -34.0688 | -6.1652
3.4367 | -34.069 | -6.1752
3.4428 | -34.0693 | -6.1846
3.4489 | -34.0693 | -6.1956
3.455 | -34.0693 | -6.206
3.4611 | -34.0693 | -6.2167
3.4672 | -34.0694 | -6.2278
3.4733 | -34.0695 | -6.2377
3.4794 | -34.0695 | -6.248
3.4855 | -34.0698 | -6.2579
3.4916 | -34.0699 | -6.268
3.4977 | -34.0699 | -6.2785
3.5038 | -34.0701 | -6.2884
3.5099 | -34.0702 | -6.2984
3.5161 | -34.0704 | -6.3082
3.5222 | -34.0706 | -6.3186
3.5283 | -34.0706 | -6.3286
3.5344 | -34.0707 | -6.3393
3.5405 | -34.0709 | -6.3494
3.5466 | -34.071 | -6.359
3.5527 | -34.0712 | -6.3696
3.5588 | -34.0714 | -6.3795
3.5649 | -34.0715 | -6.3897
3.571 | -34.0717 | -6.4
3.5771 | -34.0719 | -6.4105
3.5832 | -34.072 | -6.4207
3.5893 | -34.0722 | -6.4306
3.5954 | -34.0724 | -6.441
3.6015 | -34.0725 | -6.4512
3.6076 | -34.0728 | -6.461
3.6137 | -34.0729 | -6.4714
3.6198 | -34.0732 | -6.4806
3.6259 | -34.0733 | -6.4912
3.632 | -34.0736 | -6.5018
3.6381 | -34.0738 | -6.5116
3.6442 | -34.0741 | -6.5221
3.6503 | -34.0743 | -6.5315
3.6565 | -34.0746 | -6.5431
3.6626 | -34.0748 | -6.554
3.6687 | -34.075 | -6.5645
3.6748 | -34.0753 | -6.5746
3.6809 | -34.0756 | -6.585
3.687 | -34.0759 | -6.5959
3.6931 | -34.0762 | -6.6062
3.6992 | -34.0764 | -6.6168
3.7053 | -34.0767 | -6.6269
3.7114 | -34.0769 | -6.6382
3.7175 | -34.0772 | -6.648
3.7236 | -34.0774 | -6.6572
3.7297 | -34.0775 | -6.6678
3.7358 | -34.0778 | -6.6792
3.7419 | -34.078 | -6.6906
3.748 | -34.0781 | -6.7014
3.7541 | -34.0783 | -6.7124
3.7602 | -34.0786 | -6.7232
3.7663 | -34.0788 | -6.7335
3.7724 | -34.0791 | -6.7444
3.7785 | -34.0794 | -6.7549
3.7846 | -34.0796 | -6.7657
3.7907 | -34.08 | -6.7763
3.7969 | -34.0801 | -6.7877
3.803 | -34.0805 | -6.798
3.8091 | -34.0808 | -6.8084
3.8152 | -34.081 | -6.8195
3.8213 | -34.0812 | -6.8309
3.8274 | -34.0815 | -6.8416
3.8335 | -34.0818 | -6.8529
3.8396 | -34.0822 | -6.8643
3.8457 | -34.0825 | -6.8743
3.8518 | -34.0829 | -6.8852
3.8579 | -34.0832 | -6.8957
3.864 | -34.0835 | -6.9062
3.8701 | -34.0838 | -6.9167
3.8762 | -34.0841 | -6.928
3.8823 | -34.0844 | -6.9384
3.8884 | -34.0847 | -6.9486
3.8945 | -34.0851 | -6.9595
3.9006 | -34.0854 | -6.9707
3.9067 | -34.0857 | -6.982
3.9128 | -34.0861 | -6.9931
3.9189 | -34.0864 | -7.0037
3.925 | -34.0867 | -7.0144
3.9311 | -34.087 | -7.0246
3.9372 | -34.0873 | -7.0358
3.9434 | -34.0877 | -7.0462
3.9495 | -34.0879 | -7.0575
3.9556 | -34.0883 | -7.0687
3.9617 | -34.0886 | -7.0798
3.9678 | -34.0888 | -7.0905
3.9739 | -34.0892 | -7.1017
3.98 | -34.0896 | -7.1134
3.9861 | -34.0898 | -7.1237
3.9922 | -34.0901 | -7.1352
3.9983 | -34.0904 | -7.1455
4.0044 | -34.0906 | -7.1554
4.0105 | -34.091 | -7.1661
4.0166 | -34.0914 | -7.177
4.0227 | -34.0917 | -7.1872
4.0288 | -34.0919 | -7.1982
4.0349 | -34.0923 | -7.2081
4.041 | -34.0926 | -7.219
4.0471 | -34.0928 | -7.2291
4.0532 | -34.0932 | -7.2404
4.0593 | -34.0934 | -7.2517
4.0654 | -34.0937 | -7.2626
4.0715 | -34.094 | -7.2735
4.0776 | -34.0943 | -7.2837
4.0838 | -34.0946 | -7.2946
4.0899 | -34.0949 | -7.306
4.096 | -34.0952 | -7.3163
4.1021 | -34.0955 | -7.3272
4.1082 | -34.0958 | -7.3376
4.1143 | -34.0961 | -7.3479
4.1204 | -34.0965 | -7.3587
4.1265 | -34.0969 | -7.3693
4.1326 | -34.0972 | -7.3798
4.1387 | -34.0975 | -7.391
4.1448 | -34.0979 | -7.4019
4.1509 | -34.0981 | -7.4126
4.157 | -34.0984 | -7.4234
4.1631 | -34.0986 | -7.4344
4.1692 | -34.099 | -7.4452
4.1753 | -34.0993 | -7.4561
4.1814 | -34.0996 | -7.4662
4.1875 | -34.0998 | -7.4774
4.1936 | -34.1001 | -7.4886
4.1997 | -34.1004 | -7.4991
4.2058 | -34.1007 | -7.5104
4.2119 | -34.1009 | -7.5203
4.218 | -34.1012 | -7.5307
4.2241 | -34.1016 | -7.5405
4.2303 | -34.1019 | -7.5519
4.2364 | -34.1023 | -7.5634
4.2425 | -34.1026 | -7.5739
4.2486 | -34.1029 | -7.585
4.2547 | -34.1031 | -7.5946
4.2608 | -34.1033 | -7.6045
4.2669 | -34.1036 | -7.6153
4.273 | -34.104 | -7.626
4.2791 | -34.1042 | -7.6369
4.2852 | -34.1044 | -7.6471
4.2913 | -34.1046 | -7.6567
4.2974 | -34.1048 | -7.6675
4.3035 | -34.105 | -7.6781
4.3096 | -34.1052 | -7.6885
4.3157 | -34.1055 | -7.6987
4.3218 | -34.1057 | -7.7089
4.3279 | -34.106 | -7.7194
4.334 | -34.1062 | -7.7294
4.3401 | -34.1065 | -7.7398
4.3462 | -34.1067 | -7.7506
4.3523 | -34.107 | -7.7608
4.3584 | -34.1072 | -7.7698
4.3645 | -34.1074 | -7.7797
4.3707 | -34.1077 | -7.7894
4.3768 | -34.1079 | -7.8
4.3829 | -34.108 | -7.8102
4.389 | -34.1083 | -7.8201
4.3951 | -34.1086 | -7.83
4.4012 | -34.1088 | -7.8397
4.4073 | -34.1091 | -7.8502
4.4134 | -34.1093 | -7.8607
4.4195 | -34.1096 | -7.8705
4.4256 | -34.1098 | -7.8804
4.4317 | -34.1101 | -7.8901
4.4378 | -34.1103 | -7.9003
4.4439 | -34.1105 | -7.9102
4.45 | -34.1107 | -7.9205
4.4561 | -34.1108 | -7.9306
4.4622 | -34.111 | -7.9405
4.4683 | -34.1112 | -7.9499
4.4744 | -34.1114 | -7.9601
4.4805 | -34.1117 | -7.9698
4.4866 | -34.1119 | -7.9793
4.4927 | -34.1121 | -7.9896
4.4988 | -34.1123 | -7.9992
4.5049 | -34.1125 | -8.009
4.511 | -34.1128 | -8.0185
4.5172 | -34.1131 | -8.0286
4.5233 | -34.1133 | -8.0391
4.5294 | -34.1135 | -8.0489
4.5355 | -34.1138 | -8.0585
4.5416 | -34.114 | -8.0679
4.5477 | -34.1142 | -8.0775
4.5538 | -34.1144 | -8.0868
4.5599 | -34.1147 | -8.0966
4.566 | -34.1149 | -8.1066
4.5721 | -34.1152 | -8.1161
4.5782 | -34.1155 | -8.1259
4.5843 | -34.1157 | -8.1357
4.5904 | -34.116 | -8.1452
4.5965 | -34.1163 | -8.1543
4.6026 | -34.1165 | -8.1648
4.6087 | -34.1167 | -8.1742
4.6148 | -34.117 | -8.1849
4.6209 | -34.1174 | -8.1936
4.627 | -34.1177 | -8.2037
4.6331 | -34.1178 | -8.213
4.6392 | -34.1181 | -8.2224
4.6453 | -34.1183 | -8.232
4.6514 | -34.1186 | -8.2418
4.6576 | -34.1186 | -8.2522
4.6637 | -34.1189 | -8.2604
4.6698 | -34.1192 | -8.2694
4.6759 | -34.1193 | -8.2791
4.682 | -34.1195 | -8.2885
4.6881 | -34.1198 | -8.2976
4.6942 | -34.12 | -8.3071
4.7003 | -34.1202 | -8.3168
4.7064 | -34.1205 | -8.3264
4.7125 | -34.1208 | -8.3357
4.7186 | -34.1209 | -8.3446
4.7247 | -34.1213 | -8.3543
4.7308 | -34.1216 | -8.3639
4.7369 | -34.1218 | -8.3736
4.743 | -34.122 | -8.3838
4.7491 | -34.1222 | -8.3933
4.7552 | -34.1223 | -8.4027
4.7613 | -34.1225 | -8.4124
4.7674 | -34.1229 | -8.4213
4.7735 | -34.123 | -8.4313
4.7796 | -34.1232 | -8.4409
4.7857 | -34.1233 | -8.4506
4.7918 | -34.1235 | -8.4596
4.7979 | -34.1239 | -8.4688
4.8041 | -34.1239 | -8.4782
4.8102 | -34.1243 | -8.4877
4.8163 | -34.1245 | -8.4973
4.8224 | -34.1246 | -8.5061
4.8285 | -34.1249 | -8.5156
4.8346 | -34.1251 | -8.5254
4.8407 | -34.1254 | -8.5348
4.8468 | -34.1256 | -8.5444
4.8529 | -34.1258 | -8.554
4.859 | -34.126 | -8.5634
4.8651 | -34.1263 | -8.573
4.8712 | -34.1265 | -8.5826
4.8773 | -34.1267 | -8.5923
4.8834 | -34.127 | -8.6021
4.8895 | -34.1273 | -8.6124
4.8956 | -34.1275 | -8.6225
4.9017 | -34.1277 | -8.6322
4.9078 | -34.1279 | -8.642
4.9139 | -34.1281 | -8.6509
4.92 | -34.1283 | -8.6602
4.9261 | -34.1286 | -8.6695
4.9322 | -34.1289 | -8.6787
4.9383 | -34.1292 | -8.6879
4.9445 | -34.1295 | -8.6981
4.9506 | -34.1297 | -8.7072
4.9567 | -34.13 | -8.7167
4.9628 | -34.1303 | -8.7262
4.9689 | -34.1304 | -8.7358
4.975 | -34.1306 | -8.745
4.9811 | -34.1309 | -8.7551
4.9872 | -34.1311 | -8.765
4.9933 | -34.1314 | -8.7745
4.9994 | -34.1317 | -8.784
5.0055 | -34.1321 | -8.7937
5.0116 | -34.1324 | -8.8036
5.0177 | -34.1327 | -8.8135
5.0238 | -34.133 | -8.8237
5.0299 | -34.1334 | -8.8344
5.036 | -34.1336 | -8.844
5.0421 | -34.1338 | -8.8538
5.0482 | -34.1341 | -8.8641
5.0543 | -34.1343 | -8.8736
5.0604 | -34.1347 | -8.8838
5.0665 | -34.135 | -8.8938
5.0726 | -34.1352 | -8.9039
5.0787 | -34.1355 | -8.914
5.0848 | -34.1358 | -8.9239
5.091 | -34.1361 | -8.9339
5.0971 | -34.1364 | -8.9443
5.1032 | -34.1367 | -8.9543
5.1093 | -34.137 | -8.9642
5.1154 | -34.1372 | -8.9744
5.1215 | -34.1374 | -8.9846
5.1276 | -34.1377 | -8.9946
5.1337 | -34.1381 | -9.0044
5.1398 | -34.1383 | -9.0144
5.1459 | -34.1387 | -9.0245
5.152 | -34.139 | -9.0354
5.1581 | -34.1392 | -9.0459
5.1642 | -34.1394 | -9.0558
5.1703 | -34.1398 | -9.0659
5.1764 | -34.14 | -9.0759
5.1825 | -34.1402 | -9.0864
5.1886 | -34.1405 | -9.0964
5.1947 | -34.1408 | -9.1072
5.2008 | -34.1411 | -9.117
5.2069 | -34.1415 | -9.1279
5.213 | -34.1418 | -9.1376
5.2191 | -34.1421 | -9.1479
5.2252 | -34.1423 | -9.1581
5.2314 | -34.1426 | -9.1686
5.2375 | -34.143 | -9.1787
5.2436 | -34.1432 | -9.1889
5.2497 | -34.1435 | -9.1989
5.2558 | -34.1439 | -9.2096
5.2619 | -34.1441 | -9.2201
5.268 | -34.1444 | -9.2304
5.2741 | -34.1447 | -9.2414
5.2802 | -34.145 | -9.2519
5.2863 | -34.1453 | -9.262
5.2924 | -34.1456 | -9.273
5.2985 | -34.146 | -9.2842
5.3046 | -34.1464 | -9.2953
5.3107 | -34.1467 | -9.3058
5.3168 | -34.1471 | -9.3162
5.3229 | -34.1474 | -9.3259
5.329 | -34.1478 | -9.3357
5.3351 | -34.1482 | -9.3461
5.3412 | -34.1485 | -9.3562
5.3473 | -34.1489 | -9.367
5.3534 | -34.1492 | -9.3776
5.3595 | -34.1495 | -9.3882
5.3656 | -34.1498 | -9.3991
5.3717 | -34.1501 | -9.4093
5.3779 | -34.1504 | -9.4198
5.384 | -34.1507 | -9.4305
5.3901 | -34.1511 | -9.4408
5.3962 | -34.1513 | -9.4516
5.4023 | -34.1517 | -9.4621
5.4084 | -34.152 | -9.4724
5.4145 | -34.1523 | -9.4835
5.4206 | -34.1525 | -9.4947
5.4267 | -34.1528 | -9.5056
5.4328 | -34.1532 | -9.5156
5.4389 | -34.1535 | -9.526
5.445 | -34.1538 | -9.5369
5.4511 | -34.1541 | -9.5473
5.4572 | -34.1545 | -9.5582
5.4633 | -34.1548 | -9.5685
5.4694 | -34.1551 | -9.5789
5.4755 | -34.1554 | -9.5898
5.4816 | -34.1557 | -9.6002
5.4877 | -34.156 | -9.6109
5.4938 | -34.1563 | -9.6214
5.4999 | -34.1566 | -9.632
5.506 | -34.1569 | -9.6426
5.5121 | -34.1572 | -9.6537
5.5183 | -34.1575 | -9.6639
5.5244 | -34.1578 | -9.6744
5.5305 | -34.1581 | -9.6852
5.5366 | -34.1584 | -9.6958
5.5427 | -34.1587 | -9.7068
5.5488 | -34.159 | -9.7177
5.5549 | -34.1593 | -9.7283
5.561 | -34.1596 | -9.7393
5.5671 | -34.16 | -9.7497
5.5732 | -34.1602 | -9.7604
5.5793 | -34.1606 | -9.7715
5.5854 | -34.161 | -9.7824
5.5915 | -34.1611 | -9.793
5.5976 | -34.1615 | -9.8038
5.6037 | -34.1618 | -9.8143
5.6098 | -34.1621 | -9.8248
5.6159 | -34.1624 | -9.8358
5.622 | -34.1627 | -9.846
5.6281 | -34.163 | -9.8567
5.6342 | -34.1633 | -9.8672
5.6403 | -34.1636 | -9.8772
5.6464 | -34.164 | -9.8878
5.6525 | -34.1642 | -9.898
5.6586 | -34.1646 | -9.9081
5.6648 | -34.1649 | -9.9192
5.6709 | -34.1652 | -9.93
5.677 | -34.1655 | -9.9413
5.6831 | -34.1658 | -9.9514
5.6892 | -34.1661 | -9.9619
5.6953 | -34.1664 | -9.9725
5.7014 | -34.1667 | -9.9836
5.7075 | -34.167 | -9.9945
5.7136 | -34.1673 | -10.005
5.7197 | -34.1677 | -10.0154
5.7258 | -34.1679 | -10.0261
5.7319 | -34.1682 | -10.0366
5.738 | -34.1685 | -10.0472
5.7441 | -34.1688 | -10.0578
5.7502 | -34.1692 | -10.0682
5.7563 | -34.1695 | -10.0785
5.7624 | -34.1698 | -10.0896
5.7685 | -34.17 | -10.1
5.7746 | -34.1703 | -10.1102
5.7807 | -34.1706 | -10.1204
5.7868 | -34.1709 | -10.1308
5.7929 | -34.1712 | -10.1413
5.799 | -34.1715 | -10.1519
5.8052 | -34.1718 | -10.1627
5.8113 | -34.172 | -10.1734
5.8174 | -34.1723 | -10.1842
5.8235 | -34.1726 | -10.1946
5.8296 | -34.1729 | -10.2056
5.8357 | -34.1731 | -10.2164
5.8418 | -34.1734 | -10.227
5.8479 | -34.1737 | -10.2378
5.854 | -34.174 | -10.2485
5.8601 | -34.1742 | -10.2589
5.8662 | -34.1745 | -10.2697
5.8723 | -34.1748 | -10.2798
5.8784 | -34.1751 | -10.2904
5.8845 | -34.1754 | -10.3011
5.8906 | -34.1757 | -10.3115
5.8967 | -34.176 | -10.3219
5.9028 | -34.1763 | -10.3326
5.9089 | -34.1765 | -10.3425
5.915 | -34.1768 | -10.3532
5.9211 | -34.1772 | -10.3639
5.9272 | -34.1774 | -10.3745
5.9333 | -34.1777 | -10.3845
5.9394 | -34.178 | -10.3948
5.9455 | -34.1783 | -10.405
5.9517 | -34.1787 | -10.4156
5.9578 | -34.1789 | -10.4263
5.9639 | -34.1792 | -10.4367
5.97 | -34.1795 | -10.4472
5.9761 | -34.1798 | -10.4574
5.9822 | -34.1802 | -10.4679
5.9883 | -34.1805 | -10.4781
5.9944 | -34.1808 | -10.4885
6.0005 | -34.1812 | -10.4991
6.0066 | -34.1814 | -10.5093
6.0127 | -34.1817 | -10.5193
6.0188 | -34.182 | -10.5299
6.0249 | -34.1823 | -10.5402
6.031 | -34.1826 | -10.5507
6.0371 | -34.1829 | -10.5613
6.0432 | -34.1833 | -10.5719
6.0493 | -34.1836 | -10.5822
6.0554 | -34.184 | -10.5926
6.0615 | -34.1842 | -10.6035
6.0676 | -34.1845 | -10.6144
6.0737 | -34.1849 | -10.6249
6.0798 | -34.1851 | -10.6352
6.0859 | -34.1854 | -10.6455
6.0921 | -34.1858 | -10.6559
6.0982 | -34.1861 | -10.6666
6.1043 | -34.1865 | -10.6769
6.1104 | -34.1868 | -10.6876
6.1165 | -34.1871 | -10.6976
6.1226 | -34.1874 | -10.7075
6.1287 | -34.1878 | -10.7181
6.1348 | -34.1882 | -10.7285
6.1409 | -34.1884 | -10.7389
6.147 | -34.1888 | -10.7491
6.1531 | -34.1891 | -10.7595
6.1592 | -34.1895 | -10.77
6.1653 | -34.1898 | -10.7802
6.1714 | -34.1901 | -10.7906
6.1775 | -34.1904 | -10.8012
6.1836 | -34.1907 | -10.8114
6.1897 | -34.191 | -10.8217
6.1958 | -34.1913 | -10.8321
6.2019 | -34.1917 | -10.8426
6.208 | -34.192 | -10.8534
6.2141 | -34.1923 | -10.864
6.2202 | -34.1927 | -10.8747
6.2263 | -34.193 | -10.8849
6.2325 | -34.1934 | -10.8951
6.2386 | -34.1938 | -10.9053
6.2447 | -34.194 | -10.9157
6.2508 | -34.1943 | -10.926
6.2569 | -34.1947 | -10.9366
6.263 | -34.195 | -10.9467
6.2691 | -34.1954 | -10.9571
6.2752 | -34.1958 | -10.9673
6.2813 | -34.1961 | -10.9777
6.2874 | -34.1964 | -10.9881
6.2935 | -34.1968 | -10.9982
6.2996 | -34.1972 | -11.0088
6.3057 | -34.1975 | -11.019
6.3118 | -34.1979 | -11.0292
6.3179 | -34.1982 | -11.0396
6.324 | -34.1986 | -11.0501
6.3301 | -34.199 | -11.0602
6.3362 | -34.1993 | -11.0706
6.3423 | -34.1997 | -11.0809
6.3484 | -34.2001 | -11.0913
6.3545 | -34.2005 | -11.1013
6.3606 | -34.2008 | -11.1116
6.3667 | -34.2013 | -11.1218
6.3728 | -34.2016 | -11.1321
6.379 | -34.2019 | -11.1424
6.3851 | -34.2023 | -11.153
6.3912 | -34.2027 | -11.1631
6.3973 | -34.2031 | -11.1731
6.4034 | -34.2035 | -11.1834
6.4095 | -34.2038 | -11.194
6.4156 | -34.2041 | -11.2039
6.4217 | -34.2045 | -11.2143
6.4278 | -34.2048 | -11.2244
6.4339 | -34.2052 | -11.2346
6.44 | -34.2056 | -11.2446
6.4461 | -34.2059 | -11.2553
6.4522 | -34.2063 | -11.2651
6.4583 | -34.2066 | -11.2754
6.4644 | -34.2069 | -11.286
6.4705 | -34.2074 | -11.2966
6.4766 | -34.2077 | -11.3073
6.4827 | -34.2081 | -11.318
6.4888 | -34.2084 | -11.3283
6.4949 | -34.2088 | -11.3384
6.501 | -34.2091 | -11.3489
6.5071 | -34.2095 | -11.3594
6.5132 | -34.2099 | -11.3695
6.5194 | -34.2103 | -11.3797
6.5255 | -34.2107 | -11.3897
6.5316 | -34.211 | -11.4001
6.5377 | -34.2114 | -11.4106
6.5438 | -34.2118 | -11.4209
6.5499 | -34.2122 | -11.4311
6.556 | -34.2125 | -11.4412
6.5621 | -34.2129 | -11.4515
6.5682 | -34.2132 | -11.4615
6.5743 | -34.2135 | -11.4725
6.5804 | -34.2138 | -11.4827
6.5865 | -34.2141 | -11.4932
6.5926 | -34.2145 | -11.5036
6.5987 | -34.2149 | -11.5137
6.6048 | -34.2152 | -11.5246
6.6109 | -34.2156 | -11.5346
6.617 | -34.216 | -11.5457
6.6231 | -34.2163 | -11.5563
6.6292 | -34.2167 | -11.5664
6.6353 | -34.2171 | -11.5774
6.6414 | -34.2175 | -11.5881
6.6475 | -34.2179 | -11.5984
6.6536 | -34.2182 | -11.6084
6.6597 | -34.2187 | -11.6187
6.6659 | -34.219 | -11.6288
6.672 | -34.2194 | -11.6392
6.6781 | -34.2198 | -11.6493
6.6842 | -34.2201 | -11.6598
6.6903 | -34.2205 | -11.67
6.6964 | -34.2208 | -11.6804
6.7025 | -34.2212 | -11.6908
6.7086 | -34.2215 | -11.701
6.7147 | -34.2219 | -11.7113
6.7208 | -34.2223 | -11.722
6.7269 | -34.2226 | -11.732
6.733 | -34.2229 | -11.7417
6.7391 | -34.2234 | -11.7521
6.7452 | -34.2238 | -11.762
6.7513 | -34.2242 | -11.7721
6.7574 | -34.2245 | -11.7826
6.7635 | -34.2249 | -11.7926
6.7696 | -34.2254 | -11.803
6.7757 | -34.2257 | -11.8135
6.7818 | -34.2261 | -11.8239
6.7879 | -34.2265 | -11.8342
6.794 | -34.2268 | -11.8445
6.8001 | -34.2272 | -11.8546
6.8063 | -34.2276 | -11.8651
6.8124 | -34.2279 | -11.8753
6.8185 | -34.2283 | -11.8854
6.8246 | -34.2287 | -11.8958
6.8307 | -34.2291 | -11.906
6.8368 | -34.2295 | -11.9164
6.8429 | -34.2298 | -11.9265
6.849 | -34.2302 | -11.9367
6.8551 | -34.2305 | -11.947
6.8612 | -34.2309 | -11.9571
6.8673 | -34.2313 | -11.9672
6.8734 | -34.2317 | -11.978
6.8795 | -34.2321 | -11.9885
6.8856 | -34.2325 | -11.9983
6.8917 | -34.2328 | -12.0079
6.8978 | -34.2332 | -12.0183
6.9039 | -34.2335 | -12.0286
6.91 | -34.2339 | -12.0388
6.9161 | -34.2342 | -12.049
6.9222 | -34.2346 | -12.0593
6.9283 | -34.2349 | -12.0693
6.9344 | -34.2353 | -12.0794
6.9405 | -34.2356 | -12.0894
6.9466 | -34.236 | -12.0997
6.9528 | -34.2363 | -12.1097
6.9589 | -34.2366 | -12.1196
6.965 | -34.237 | -12.1297
6.9711 | -34.2374 | -12.1398
6.9772 | -34.2378 | -12.1497
6.9833 | -34.2381 | -12.1601
6.9894 | -34.2384 | -12.1706
6.9955 | -34.2388 | -12.1808
7.0016 | -34.2391 | -12.1914
7.0077 | -34.2395 | -12.2019
7.0138 | -34.2399 | -12.2124
7.0199 | -34.2402 | -12.2225
7.026 | -34.2405 | -12.2326
7.0321 | -34.2408 | -12.2429
7.0382 | -34.2411 | -12.2531
7.0443 | -34.2416 | -12.2634
7.0504 | -34.2419 | -12.2739
7.0565 | -34.2422 | -12.2843
7.0626 | -34.2425 | -12.2945
7.0687 | -34.2429 | -12.3046
7.0748 | -34.2432 | -12.3148
7.0809 | -34.2435 | -12.3251
7.087 | -34.2439 | -12.3353
7.0932 | -34.2442 | -12.3454
7.0993 | -34.2446 | -12.3557
7.1054 | -34.2449 | -12.3656
7.1115 | -34.2452 | -12.3758
7.1176 | -34.2454 | -12.3859
7.1237 | -34.2457 | -12.3962
7.1298 | -34.246 | -12.406
7.1359 | -34.2463 | -12.4158
7.142 | -34.2467 | -12.4262
7.1481 | -34.247 | -12.4362
7.1542 | -34.2473 | -12.4462
7.1603 | -34.2476 | -12.4562
7.1664 | -34.248 | -12.4664
7.1725 | -34.2483 | -12.4763
7.1786 | -34.2487 | -12.4865
7.1847 | -34.249 | -12.4964
7.1908 | -34.2493 | -12.5066
7.1969 | -34.2497 | -12.5165
7.203 | -34.25 | -12.5268
7.2091 | -34.2503 | -12.5368
7.2152 | -34.2506 | -12.5468
7.2213 | -34.2509 | -12.5569
7.2274 | -34.2512 | -12.5669
7.2335 | -34.2516 | -12.5771
7.2397 | -34.2519 | -12.5872
7.2458 | -34.2522 | -12.597
7.2519 | -34.2525 | -12.607
7.258 | -34.2528 | -12.6172
7.2641 | -34.2531 | -12.6271
7.2702 | -34.2535 | -12.6374
7.2763 | -34.2538 | -12.6471
7.2824 | -34.2542 | -12.657
7.2885 | -34.2546 | -12.6669
7.2946 | -34.2549 | -12.6772
7.3007 | -34.2552 | -12.6872
7.3068 | -34.2555 | -12.6974
7.3129 | -34.2558 | -12.7071
7.319 | -34.2562 | -12.7172
7.3251 | -34.2565 | -12.7271
7.3312 | -34.2569 | -12.7368
7.3373 | -34.2572 | -12.7466
7.3434 | -34.2576 | -12.7564
7.3495 | -34.2579 | -12.766
7.3556 | -34.2582 | -12.7763
7.3617 | -34.2585 | -12.7866
7.3678 | -34.2588 | -12.7968
7.3739 | -34.2592 | -12.8062
7.3801 | -34.2595 | -12.8159
7.3862 | -34.2598 | -12.8258
7.3923 | -34.2601 | -12.8356
7.3984 | -34.2605 | -12.8456
7.4045 | -34.2608 | -12.8555
7.4106 | -34.2611 | -12.8653
7.4167 | -34.2614 | -12.8754
7.4228 | -34.2617 | -12.8851
7.4289 | -34.262 | -12.8948
7.435 | -34.2624 | -12.9046
7.4411 | -34.2627 | -12.9142
7.4472 | -34.2631 | -12.9242
7.4533 | -34.2635 | -12.9343
7.4594 | -34.2637 | -12.9443
7.4655 | -34.2641 | -12.9543
7.4716 | -34.2644 | -12.9642
7.4777 | -34.2648 | -12.9737
7.4838 | -34.2651 | -12.9839
7.4899 | -34.2655 | -12.9939
7.496 | -34.2658 | -13.004
7.5021 | -34.2661 | -13.0138
7.5082 | -34.2664 | -13.0237
7.5143 | -34.2668 | -13.0335
7.5204 | -34.2671 | -13.043
7.5266 | -34.2674 | -13.053
7.5327 | -34.2677 | -13.0629
7.5388 | -34.268 | -13.0725
7.5449 | -34.2683 | -13.0823
7.551 | -34.2686 | -13.0925
7.5571 | -34.2689 | -13.1022
7.5632 | -34.2693 | -13.1119
7.5693 | -34.2696 | -13.122
7.5754 | -34.27 | -13.132
7.5815 | -34.2703 | -13.1415
7.5876 | -34.2706 | -13.1511
7.5937 | -34.2709 | -13.1607
7.5998 | -34.2712 | -13.1706
7.6059 | -34.2714 | -13.1802
7.612 | -34.2718 | -13.19
7.6181 | -34.2721 | -13.1998
7.6242 | -34.2725 | -13.2095
7.6303 | -34.2728 | -13.2189
7.6364 | -34.2732 | -13.2289
7.6425 | -34.2735 | -13.239
7.6486 | -34.2738 | -13.2488
7.6547 | -34.2742 | -13.2588
7.6608 | -34.2745 | -13.2686
7.667 | -34.2749 | -13.2779
7.6731 | -34.2751 | -13.2878
7.6792 | -34.2754 | -13.2975
7.6853 | -34.2757 | -13.3071
7.6914 | -34.276 | -13.317
7.6975 | -34.2763 | -13.3269
7.7036 | -34.2766 | -13.3368
7.7097 | -34.2769 | -13.3463
7.7158 | -34.2772 | -13.3559
7.7219 | -34.2775 | -13.3657
7.728 | -34.2778 | -13.3756
7.7341 | -34.2781 | -13.3853
7.7402 | -34.2785 | -13.3949
7.7463 | -34.2787 | -13.4047
7.7524 | -34.279 | -13.4146
7.7585 | -34.2793 | -13.4245
7.7646 | -34.2797 | -13.434
7.7707 | -34.2799 | -13.4435
7.7768 | -34.2802 | -13.453
7.7829 | -34.2805 | -13.4626
7.789 | -34.2808 | -13.4719
7.7951 | -34.2811 | -13.482
7.8012 | -34.2814 | -13.4921
7.8073 | -34.2817 | -13.5015
7.8135 | -34.282 | -13.5113
7.8196 | -34.2823 | -13.5208
7.8257 | -34.2826 | -13.5308
7.8318 | -34.283 | -13.5409
7.8379 | -34.2833 | -13.5503
7.844 | -34.2835 | -13.5596
7.8501 | -34.2837 | -13.5695
7.8562 | -34.284 | -13.5796
7.8623 | -34.2844 | -13.5894
7.8684 | -34.2846 | -13.5988
7.8745 | -34.285 | -13.6084
7.8806 | -34.2852 | -13.6175
7.8867 | -34.2856 | -13.6272
7.8928 | -34.2859 | -13.6368
7.8989 | -34.2862 | -13.6461
7.905 | -34.2866 | -13.6561
7.9111 | -34.2869 | -13.6658
7.9172 | -34.2872 | -13.6752
7.9233 | -34.2875 | -13.6849
7.9294 | -34.2878 | -13.6946
7.9355 | -34.2881 | -13.704
7.9416 | -34.2884 | -13.7135
7.9477 | -34.2887 | -13.7229
7.9539 | -34.289 | -13.7323
7.96 | -34.2893 | -13.7422
7.9661 | -34.2896 | -13.7518
7.9722 | -34.29 | -13.7617
7.9783 | -34.2902 | -13.7711
7.9844 | -34.2906 | -13.7812
7.9905 | -34.2909 | -13.7906
7.9966 | -34.2912 | -13.8001
8.0027 | -34.2915 | -13.8101
8.0088 | -34.2918 | -13.8194
8.0149 | -34.2921 | -13.8288
8.021 | -34.2924 | -13.8384
8.0271 | -34.2927 | -13.8482
8.0332 | -34.293 | -13.858
8.0393 | -34.2933 | -13.8676
8.0454 | -34.2936 | -13.8773
8.0515 | -34.2939 | -13.8868
8.0576 | -34.2942 | -13.8965
8.0637 | -34.2945 | -13.9058
8.0698 | -34.2948 | -13.9152
8.0759 | -34.2951 | -13.9246
8.082 | -34.2954 | -13.9344
8.0881 | -34.2957 | -13.9439
8.0942 | -34.296 | -13.9538
8.1004 | -34.2963 | -13.9636
8.1065 | -34.2966 | -13.9733
8.1126 | -34.297 | -13.9829
8.1187 | -34.2973 | -13.9925
8.1248 | -34.2976 | -14.0025
8.1309 | -34.2979 | -14.0122
8.137 | -34.2982 | -14.0219
8.1431 | -34.2985 | -14.0315
8.1492 | -34.2988 | -14.0412
8.1553 | -34.2991 | -14.0508
8.1614 | -34.2994 | -14.0605
8.1675 | -34.2997 | -14.07
8.1736 | -34.3001 | -14.0798
8.1797 | -34.3003 | -14.0895
8.1858 | -34.3006 | -14.099
8.1919 | -34.3009 | -14.1085
8.198 | -34.3012 | -14.1179
8.2041 | -34.3015 | -14.1278
8.2102 | -34.3019 | -14.1374
8.2163 | -34.3022 | -14.1471
8.2224 | -34.3025 | -14.1569
8.2285 | -34.3028 | -14.1666
8.2346 | -34.3032 | -14.1762
8.2408 | -34.3034 | -14.186
8.2469 | -34.3037 | -14.1954
8.253 | -34.304 | -14.2051
8.2591 | -34.3043 | -14.2143
8.2652 | -34.3047 | -14.2236
8.2713 | -34.3051 | -14.2332
8.2774 | -34.3053 | -14.2428
8.2835 | -34.3057 | -14.2526
8.2896 | -34.306 | -14.2618
8.2957 | -34.3063 | -14.2714
8.3018 | -34.3066 | -14.2809
8.3079 | -34.3069 | -14.2905
8.314 | -34.3072 | -14.3
8.3201 | -34.3076 | -14.3095
8.3262 | -34.3079 | -14.319
8.3323 | -34.3082 | -14.3285
8.3384 | -34.3085 | -14.3384
8.3445 | -34.3088 | -14.3477
8.3506 | -34.3091 | -14.3574
8.3567 | -34.3094 | -14.3672
8.3628 | -34.3097 | -14.3768
8.3689 | -34.3101 | -14.3864
8.375 | -34.3103 | -14.396
8.3812 | -34.3107 | -14.4058
8.3873 | -34.311 | -14.4154
8.3934 | -34.3113 | -14.4249
8.3995 | -34.3116 | -14.4347
8.4056 | -34.312 | -14.4443
8.4117 | -34.3123 | -14.4541
8.4178 | -34.3126 | -14.4633
8.4239 | -34.3129 | -14.4728
8.43 | -34.3133 | -14.4823
8.4361 | -34.3136 | -14.4923
8.4422 | -34.3139 | -14.5018
8.4483 | -34.3142 | -14.5114
8.4544 | -34.3145 | -14.5213
8.4605 | -34.3148 | -14.5309
8.4666 | -34.3151 | -14.5405
8.4727 | -34.3154 | -14.55
8.4788 | -34.3157 | -14.5594
8.4849 | -34.3161 | -14.5693
8.491 | -34.3164 | -14.5788
8.4971 | -34.3167 | -14.5886
8.5032 | -34.317 | -14.5983
8.5093 | -34.3173 | -14.6079
8.5154 | -34.3176 | -14.6172
8.5215 | -34.3179 | -14.627
8.5277 | -34.3182 | -14.6367
8.5338 | -34.3185 | -14.646
8.5399 | -34.3188 | -14.656
8.546 | -34.3191 | -14.6652
8.5521 | -34.3195 | -14.675
8.5582 | -34.3198 | -14.6847
8.5643 | -34.3201 | -14.6943
8.5704 | -34.3204 | -14.7038
8.5765 | -34.3207 | -14.7132
8.5826 | -34.3211 | -14.723
8.5887 | -34.3214 | -14.7325
8.5948 | -34.3216 | -14.7417
8.6009 | -34.322 | -14.7514
8.607 | -34.3223 | -14.7611
8.6131 | -34.3226 | -14.771
8.6192 | -34.3229 | -14.7802
8.6253 | -34.3232 | -14.7901
8.6314 | -34.3236 | -14.7997
8.6375 | -34.3239 | -14.8095
8.6436 | -34.3243 | -14.8191
8.6497 | -34.3246 | -14.8288
8.6558 | -34.3249 | -14.8383
8.6619 | -34.3252 | -14.8481
8.6681 | -34.3255 | -14.8574
8.6742 | -34.3258 | -14.8672
8.6803 | -34.3262 | -14.8772
8.6864 | -34.3265 | -14.8872
8.6925 | -34.3268 | -14.897
8.6986 | -34.3271 | -14.9069
8.7047 | -34.3275 | -14.9164
8.7108 | -34.3278 | -14.9261
8.7169 | -34.3281 | -14.9355
8.723 | -34.3285 | -14.9452
8.7291 | -34.3288 | -14.9549
8.7352 | -34.3292 | -14.9647
8.7413 | -34.3295 | -14.9742
8.7474 | -34.3298 | -14.9837
8.7535 | -34.3302 | -14.9934
8.7596 | -34.3306 | -15.0035
8.7657 | -34.3309 | -15.013
8.7718 | -34.3312 | -15.0223
8.7779 | -34.3315 | -15.0324
8.784 | -34.3318 | -15.0418
8.7901 | -34.332 | -15.0518
8.7962 | -34.3324 | -15.0617
8.8023 | -34.3328 | -15.071
8.8084 | -34.3331 | -15.0809
8.8146 | -34.3334 | -15.0904
8.8207 | -34.3338 | -15.1001
8.8268 | -34.3341 | -15.1098
8.8329 | -34.3345 | -15.1196
8.839 | -34.3348 | -15.1292
8.8451 | -34.3352 | -15.1388
8.8512 | -34.3355 | -15.1486
8.8573 | -34.3358 | -15.1581
8.8634 | -34.3362 | -15.1677
8.8695 | -34.3365 | -15.1773
8.8756 | -34.3369 | -15.1867
8.8817 | -34.3372 | -15.1967
8.8878 | -34.3375 | -15.2064
8.8939 | -34.3378 | -15.2157
8.9 | -34.3381 | -15.225
8.9061 | -34.3385 | -15.2347
8.9122 | -34.3388 | -15.2442
8.9183 | -34.3391 | -15.2537
8.9244 | -34.3395 | -15.2629
8.9305 | -34.3398 | -15.2727
8.9366 | -34.3402 | -15.2822
8.9427 | -34.3405 | -15.2919
8.9488 | -34.3409 | -15.3015
8.955 | -34.3412 | -15.3111
8.9611 | -34.3415 | -15.3207
8.9672 | -34.3419 | -15.3306
8.9733 | -34.3422 | -15.3402
8.9794 | -34.3425 | -15.35
8.9855 | -34.3429 | -15.3595
8.9916 | -34.3432 | -15.3688
8.9977 | -34.3436 | -15.3784
9.0038 | -34.344 | -15.3882
9.0099 | -34.3443 | -15.398
9.016 | -34.3447 | -15.4075
9.0221 | -34.345 | -15.4176
9.0282 | -34.3454 | -15.4275
9.0343 | -34.3457 | -15.4374
9.0404 | -34.3461 | -15.4468
9.0465 | -34.3465 | -15.4567
9.0526 | -34.3468 | -15.4663
9.0587 | -34.3471 | -15.4758
9.0648 | -34.3475 | -15.4854
9.0709 | -34.3479 | -15.4949
9.077 | -34.3482 | -15.505
9.0831 | -34.3485 | -15.5145
9.0892 | -34.3488 | -15.5241
9.0953 | -34.3491 | -15.5337
9.1015 | -34.3495 | -15.5434
9.1076 | -34.3499 | -15.5532
9.1137 | -34.3502 | -15.5628
9.1198 | -34.3506 | -15.572
9.1259 | -34.351 | -15.5817
9.132 | -34.3513 | -15.5915
9.1381 | -34.3517 | -15.6011
9.1442 | -34.3521 | -15.6104
9.1503 | -34.3524 | -15.6202
9.1564 | -34.3528 | -15.6298
9.1625 | -34.3531 | -15.6395
9.1686 | -34.3535 | -15.6491
9.1747 | -34.3538 | -15.6588
9.1808 | -34.3542 | -15.6686
9.1869 | -34.3545 | -15.6785
9.193 | -34.3549 | -15.6883
9.1991 | -34.3552 | -15.6976
9.2052 | -34.3555 | -15.7071
9.2113 | -34.3559 | -15.717
9.2174 | -34.3563 | -15.7271
9.2235 | -34.3566 | -15.7364
9.2296 | -34.357 | -15.7462
9.2357 | -34.3574 | -15.7557
9.2419 | -34.3577 | -15.7654
9.248 | -34.3581 | -15.7751
9.2541 | -34.3584 | -15.7849
9.2602 | -34.3588 | -15.7945
9.2663 | -34.3591 | -15.8042
9.2724 | -34.3595 | -15.8139
9.2785 | -34.3598 | -15.8236
9.2846 | -34.3602 | -15.8333
9.2907 | -34.3605 | -15.8429
9.2968 | -34.3609 | -15.8525
9.3029 | -34.3612 | -15.8622
9.309 | -34.3616 | -15.8718
9.3151 | -34.3619 | -15.8814
9.3212 | -34.3623 | -15.8912
9.3273 | -34.3627 | -15.9007
9.3334 | -34.3631 | -15.9103
9.3395 | -34.3635 | -15.92
9.3456 | -34.3639 | -15.9295
9.3517 | -34.3643 | -15.9387
9.3578 | -34.3647 | -15.9481
9.3639 | -34.365 | -15.9581
9.37 | -34.3654 | -15.9677
9.3761 | -34.3658 | -15.9771
9.3822 | -34.3661 | -15.9865
9.3884 | -34.3665 | -15.9962
9.3945 | -34.3669 | -16.0058
9.4006 | -34.3672 | -16.0155
9.4067 | -34.3676 | -16.0253
9.4128 | -34.368 | -16.0347
9.4189 | -34.3684 | -16.0443
9.425 | -34.3687 | -16.0538
9.4311 | -34.3691 | -16.0634
9.4372 | -34.3694 | -16.073
9.4433 | -34.3698 | -16.0828
9.4494 | -34.3701 | -16.0925
9.4555 | -34.3705 | -16.102
9.4616 | -34.3709 | -16.1115
9.4677 | -34.3713 | -16.121
9.4738 | -34.3717 | -16.131
9.4799 | -34.372 | -16.1405
9.486 | -34.3723 | -16.1504
9.4921 | -34.3727 | -16.1605
9.4982 | -34.3731 | -16.1698
9.5043 | -34.3734 | -16.1791
9.5104 | -34.3738 | -16.1887
9.5165 | -34.3742 | -16.1983
9.5226 | -34.3746 | -16.2079
9.5288 | -34.375 | -16.2175
9.5349 | -34.3754 | -16.2271
9.541 | -34.3757 | -16.2368
9.5471 | -34.3761 | -16.2466
9.5532 | -34.3765 | -16.2561
9.5593 | -34.3769 | -16.2655
9.5654 | -34.3773 | -16.2751
9.5715 | -34.3776 | -16.285
9.5776 | -34.378 | -16.2948
9.5837 | -34.3784 | -16.3042
9.5898 | -34.3788 | -16.3139
9.5959 | -34.3791 | -16.3233
9.602 | -34.3795 | -16.3331
9.6081 | -34.3798 | -16.3429
9.6142 | -34.3802 | -16.3524
9.6203 | -34.3806 | -16.3621
9.6264 | -34.3809 | -16.3714
9.6325 | -34.3813 | -16.3811
9.6386 | -34.3817 | -16.3906
9.6447 | -34.382 | -16.4001
9.6508 | -34.3824 | -16.4099
9.6569 | -34.3828 | -16.4194
9.663 | -34.3832 | -16.429
9.6691 | -34.3836 | -16.4385
9.6753 | -34.384 | -16.4481
9.6814 | -34.3844 | -16.4576
9.6875 | -34.3848 | -16.4671
9.6936 | -34.3852 | -16.4767
9.6997 | -34.3856 | -16.4864
9.7058 | -34.386 | -16.4959
9.7119 | -34.3863 | -16.5053
9.718 | -34.3867 | -16.5151
9.7241 | -34.387 | -16.5247
9.7302 | -34.3874 | -16.5344
9.7363 | -34.3878 | -16.5438
9.7424 | -34.3882 | -16.5533
9.7485 | -34.3886 | -16.5628
9.7546 | -34.389 | -16.5723
9.7607 | -34.3893 | -16.5819
9.7668 | -34.3897 | -16.5918
9.7729 | -34.3901 | -16.6012
9.779 | -34.3906 | -16.6108
9.7851 | -34.3909 | -16.6204
9.7912 | -34.3912 | -16.6305
9.7973 | -34.3915 | -16.6402
9.8034 | -34.392 | -16.6501
9.8095 | -34.3923 | -16.6594
9.8157 | -34.3927 | -16.6689
9.8218 | -34.3931 | -16.678
9.8279 | -34.3935 | -16.6878
9.834 | -34.3939 | -16.6972
9.8401 | -34.3943 | -16.7066
9.8462 | -34.3947 | -16.716
9.8523 | -34.3951 | -16.7255
9.8584 | -34.3955 | -16.7353
9.8645 | -34.3959 | -16.7448
9.8706 | -34.3963 | -16.7542
9.8767 | -34.3966 | -16.7634
9.8828 | -34.3971 | -16.7727
9.8889 | -34.3975 | -16.7827
9.895 | -34.3978 | -16.7918
9.9011 | -34.3982 | -16.8009
9.9072 | -34.3986 | -16.8101
9.9133 | -34.399 | -16.8195
9.9194 | -34.3994 | -16.8289
9.9255 | -34.3998 | -16.8386
9.9316 | -34.4002 | -16.8479
9.9377 | -34.4006 | -16.8573
9.9438 | -34.4011 | -16.8667
9.9499 | -34.4014 | -16.8761
9.956 | -34.4019 | -16.8853
9.9622 | -34.4022 | -16.8947
9.9683 | -34.4026 | -16.9041
9.9744 | -34.403 | -16.9135
9.9805 | -34.4034 | -16.923
9.9866 | -34.4038 | -16.9321
9.9927 | -34.4043 | -16.9419
9.9988 | -34.4047 | -16.951
10.0049 | -34.405 | -16.9604
10.011 | -34.4055 | -16.9698
10.0171 | -34.4058 | -16.9793
10.0232 | -34.4063 | -16.9885
10.0293 | -34.4067 | -16.9976
10.0354 | -34.4071 | -17.0067
10.0415 | -34.4075 | -17.0161
10.0476 | -34.4079 | -17.0258
10.0537 | -34.4083 | -17.0354
10.0598 | -34.4087 | -17.045
10.0659 | -34.4091 | -17.0543
10.072 | -34.4095 | -17.0638
10.0781 | -34.4099 | -17.0733
10.0842 | -34.4103 | -17.0828
10.0903 | -34.4107 | -17.0925
10.0964 | -34.4112 | -17.1018
10.1026 | -34.4116 | -17.1112
10.1087 | -34.412 | -17.1207
10.1148 | -34.4125 | -17.1302
10.1209 | -34.4129 | -17.1396
10.127 | -34.4133 | -17.1488
10.1331 | -34.4137 | -17.1582
10.1392 | -34.4141 | -17.1676
10.1453 | -34.4145 | -17.1771
10.1514 | -34.4149 | -17.1862
10.1575 | -34.4153 | -17.1954
10.1636 | -34.4157 | -17.2042
10.1697 | -34.4161 | -17.2135
10.1758 | -34.4165 | -17.2229
10.1819 | -34.417 | -17.2324
10.188 | -34.4174 | -17.2416
10.1941 | -34.4177 | -17.251
10.2002 | -34.4181 | -17.2598
10.2063 | -34.4185 | -17.2694
10.2124 | -34.4189 | -17.2788
10.2185 | -34.4193 | -17.288
10.2246 | -34.4197 | -17.2974
10.2307 | -34.4201 | -17.3069
10.2368 | -34.4206 | -17.3156
10.2429 | -34.421 | -17.3246
10.2491 | -34.4214 | -17.3337
10.2552 | -34.4218 | -17.3434
10.2613 | -34.4223 | -17.3529
10.2674 | -34.4226 | -17.3623
10.2735 | -34.423 | -17.3717
10.2796 | -34.4234 | -17.3812
10.2857 | -34.4238 | -17.3904
10.2918 | -34.4242 | -17.3997
10.2979 | -34.4247 | -17.4092
10.304 | -34.4251 | -17.4188
10.3101 | -34.4256 | -17.4281
10.3162 | -34.426 | -17.4373
10.3223 | -34.4264 | -17.4468
10.3284 | -34.4268 | -17.456
10.3345 | -34.4272 | -17.4652
10.3406 | -34.4276 | -17.4744
10.3467 | -34.428 | -17.4837
10.3528 | -34.4284 | -17.4929
10.3589 | -34.4288 | -17.5024
10.365 | -34.4292 | -17.5118
10.3711 | -34.4296 | -17.5213
10.3772 | -34.4299 | -17.5305
10.3833 | -34.4303 | -17.5398
10.3895 | -34.4307 | -17.5489
10.3956 | -34.4312 | -17.5579
10.4017 | -34.4315 | -17.5672
10.4078 | -34.432 | -17.5766
10.4139 | -34.4324 | -17.5858
10.42 | -34.4328 | -17.595
10.4261 | -34.4333 | -17.6042
10.4322 | -34.4337 | -17.6133
10.4383 | -34.4341 | -17.6227
10.4444 | -34.4345 | -17.6316
10.4505 | -34.435 | -17.6409
10.4566 | -34.4354 | -17.6501
10.4627 | -34.4358 | -17.6595
10.4688 | -34.4362 | -17.6684
10.4749 | -34.4366 | -17.6774
10.481 | -34.4371 | -17.6866
10.4871 | -34.4375 | -17.696
10.4932 | -34.4379 | -17.7053
10.4993 | -34.4384 | -17.7144
10.5054 | -34.4388 | -17.7237
10.5115 | -34.4392 | -17.7329
10.5176 | -34.4396 | -17.7419
10.5237 | -34.44 | -17.7508
10.5298 | -34.4404 | -17.7599
10.536 | -34.4409 | -17.7693
10.5421 | -34.4413 | -17.7782
10.5482 | -34.4417 | -17.7877
10.5543 | -34.4422 | -17.7969
10.5604 | -34.4426 | -17.8058
10.5665 | -34.443 | -17.815
10.5726 | -34.4434 | -17.824
10.5787 | -34.4438 | -17.8334
10.5848 | -34.4443 | -17.8424
10.5909 | -34.4447 | -17.8515
10.597 | -34.4451 | -17.8607
10.6031 | -34.4455 | -17.8699
10.6092 | -34.446 | -17.8793
10.6153 | -34.4463 | -17.8887
10.6214 | -34.4468 | -17.8977
10.6275 | -34.4473 | -17.9069
10.6336 | -34.4477 | -17.9164
10.6397 | -34.4481 | -17.9254
10.6458 | -34.4485 | -17.9344
10.6519 | -34.4488 | -17.9434
10.658 | -34.4493 | -17.9523
10.6641 | -34.4497 | -17.9614
10.6702 | -34.4502 | -17.9708
10.6764 | -34.4506 | -17.9798
10.6825 | -34.4511 | -17.9892
10.6886 | -34.4515 | -17.9985
10.6947 | -34.4519 | -18.0074
10.7008 | -34.4523 | -18.0164
10.7069 | -34.4528 | -18.0256
10.713 | -34.4532 | -18.0347
10.7191 | -34.4537 | -18.044
10.7252 | -34.4541 | -18.053
10.7313 | -34.4546 | -18.0622
10.7374 | -34.455 | -18.0711
10.7435 | -34.4554 | -18.0805
10.7496 | -34.4558 | -18.0896
10.7557 | -34.4562 | -18.0986
10.7618 | -34.4567 | -18.1079
10.7679 | -34.4571 | -18.1169
10.774 | -34.4575 | -18.126
10.7801 | -34.4579 | -18.1355
10.7862 | -34.4584 | -18.1447
10.7923 | -34.4589 | -18.1538
10.7984 | -34.4593 | -18.1628
10.8045 | -34.4597 | -18.1722
10.8106 | -34.4601 | -18.1813
10.8168 | -34.4606 | -18.1905
10.8229 | -34.461 | -18.1997
10.829 | -34.4614 | -18.209
10.8351 | -34.4618 | -18.2182
10.8412 | -34.4622 | -18.2273
10.8473 | -34.4626 | -18.2363
10.8534 | -34.463 | -18.2456
10.8595 | -34.4634 | -18.2545
10.8656 | -34.4639 | -18.2635
10.8717 | -34.4644 | -18.2723
10.8778 | -34.4648 | -18.2812
10.8839 | -34.4652 | -18.2902
10.89 | -34.4656 | -18.2991
10.8961 | -34.466 | -18.3082
10.9022 | -34.4665 | -18.3173
10.9083 | -34.4669 | -18.3267
10.9144 | -34.4674 | -18.3358
10.9205 | -34.4679 | -18.3449
10.9266 | -34.4683 | -18.3543
10.9327 | -34.4688 | -18.3634
10.9388 | -34.4692 | -18.3725
10.9449 | -34.4696 | -18.3816
10.951 | -34.47 | -18.3907
10.9571 | -34.4704 | -18.3998
10.9633 | -34.4709 | -18.4092
10.9694 | -34.4714 | -18.4185
10.9755 | -34.4718 | -18.4275
10.9816 | -34.4722 | -18.4366
10.9877 | -34.4727 | -18.4458
10.9938 | -34.4731 | -18.4549
10.9999 | -34.4735 | -18.464
11.006 | -34.474 | -18.473
11.0121 | -34.4744 | -18.4825
11.0182 | -34.4748 | -18.4912
11.0243 | -34.4753 | -18.5003
11.0304 | -34.4757 | -18.5093
11.0365 | -34.4761 | -18.5184
11.0426 | -34.4765 | -18.5272
11.0487 | -34.477 | -18.5362
11.0548 | -34.4774 | -18.5453
11.0609 | -34.4779 | -18.5546
11.067 | -34.4783 | -18.5639
11.0731 | -34.4787 | -18.573
11.0792 | -34.4791 | -18.5821
11.0853 | -34.4796 | -18.5909
11.0914 | -34.48 | -18.5998
11.0975 | -34.4805 | -18.6092
11.1037 | -34.4809 | -18.6181
11.1098 | -34.4813 | -18.6272
11.1159 | -34.4818 | -18.6365
11.122 | -34.4823 | -18.6459
11.1281 | -34.4826 | -18.6547
11.1342 | -34.4831 | -18.6635
11.1403 | -34.4835 | -18.6727
11.1464 | -34.484 | -18.6817
11.1525 | -34.4844 | -18.691
11.1586 | -34.4848 | -18.7005
11.1647 | -34.4853 | -18.7095
11.1708 | -34.4857 | -18.7184
11.1769 | -34.4861 | -18.727
11.183 | -34.4866 | -18.736
11.1891 | -34.487 | -18.7454
11.1952 | -34.4874 | -18.7547
11.2013 | -34.4878 | -18.7637
11.2074 | -34.4883 | -18.7726
11.2135 | -34.4886 | -18.7815
11.2196 | -34.4891 | -18.7905
11.2257 | -34.4894 | -18.7994
11.2318 | -34.4898 | -18.8087
11.2379 | -34.4902 | -18.8175
11.244 | -34.4906 | -18.8264
11.2502 | -34.491 | -18.8354
11.2563 | -34.4915 | -18.8444
11.2624 | -34.4919 | -18.8536
11.2685 | -34.4923 | -18.8626
11.2746 | -34.4927 | -18.8714
11.2807 | -34.4931 | -18.8808
11.2868 | -34.4935 | -18.89
11.2929 | -34.494 | -18.8989
11.299 | -34.4944 | -18.9081
11.3051 | -34.4948 | -18.9169
11.3112 | -34.4952 | -18.9259
11.3173 | -34.4957 | -18.9352
11.3234 | -34.4961 | -18.9444
11.3295 | -34.4965 | -18.9531
11.3356 | -34.497 | -18.9619
11.3417 | -34.4974 | -18.9712
11.3478 | -34.4978 | -18.9801
11.3539 | -34.4982 | -18.989
11.36 | -34.4986 | -18.9979
11.3661 | -34.4991 | -19.0068
11.3722 | -34.4995 | -19.0159
11.3783 | -34.4999 | -19.0247
11.3844 | -34.5003 | -19.0338
11.3906 | -34.5007 | -19.0431
11.3967 | -34.5011 | -19.0522
11.4028 | -34.5015 | -19.0612
11.4089 | -34.502 | -19.0701
11.415 | -34.5024 | -19.0791
11.4211 | -34.5028 | -19.0878
11.4272 | -34.5032 | -19.0969
11.4333 | -34.5036 | -19.1061
11.4394 | -34.5041 | -19.1154
11.4455 | -34.5045 | -19.1246
11.4516 | -34.5049 | -19.1334
11.4577 | -34.5053 | -19.1427
11.4638 | -34.5058 | -19.1513
11.4699 | -34.5062 | -19.1607
11.476 | -34.5066 | -19.17
11.4821 | -34.5071 | -19.1788
11.4882 | -34.5075 | -19.188
11.4943 | -34.5079 | -19.1968
11.5004 | -34.5084 | -19.2059
11.5065 | -34.5088 | -19.215
11.5126 | -34.5092 | -19.2238
11.5187 | -34.5097 | -19.2328
11.5248 | -34.5101 | -19.2417
11.5309 | -34.5105 | -19.2507
11.5371 | -34.511 | -19.2597
11.5432 | -34.5113 | -19.2689
11.5493 | -34.5117 | -19.2779
11.5554 | -34.5122 | -19.2868
11.5615 | -34.5126 | -19.2959
11.5676 | -34.5131 | -19.305
11.5737 | -34.5135 | -19.3138
11.5798 | -34.5139 | -19.3227
11.5859 | -34.5144 | -19.3316
11.592 | -34.5148 | -19.3407
11.5981 | -34.5152 | -19.3497
11.6042 | -34.5156 | -19.3587
11.6103 | -34.5161 | -19.3677
11.6164 | -34.5165 | -19.3767
11.6225 | -34.5169 | -19.3856
11.6286 | -34.5173 | -19.3948
11.6347 | -34.5178 | -19.404
11.6408 | -34.5182 | -19.413
11.6469 | -34.5187 | -19.4222
11.653 | -34.5191 | -19.431
11.6591 | -34.5196 | -19.44
11.6652 | -34.52 | -19.4492
11.6713 | -34.5204 | -19.4582
11.6775 | -34.5208 | -19.467
11.6836 | -34.5212 | -19.4762
11.6897 | -34.5217 | -19.4849
11.6958 | -34.5221 | -19.4941
11.7019 | -34.5226 | -19.503
11.708 | -34.523 | -19.5123
11.7141 | -34.5235 | -19.5213
11.7202 | -34.5239 | -19.53
11.7263 | -34.5243 | -19.5391
11.7324 | -34.5248 | -19.5481
11.7385 | -34.5252 | -19.5573
11.7446 | -34.5256 | -19.5666
11.7507 | -34.526 | -19.5755
11.7568 | -34.5265 | -19.5846
11.7629 | -34.5269 | -19.5936
11.769 | -34.5274 | -19.6026
11.7751 | -34.5278 | -19.6113
11.7812 | -34.5282 | -19.6202
11.7873 | -34.5286 | -19.6295
11.7934 | -34.5291 | -19.6384
11.7995 | -34.5295 | -19.6471
11.8056 | -34.5299 | -19.6566
11.8117 | -34.5303 | -19.6656
11.8178 | -34.5308 | -19.6745
11.824 | -34.5313 | -19.6833
11.8301 | -34.5317 | -19.692
11.8362 | -34.5321 | -19.7007
11.8423 | -34.5325 | -19.7098
11.8484 | -34.533 | -19.7187
11.8545 | -34.5334 | -19.728
11.8606 | -34.5339 | -19.7371
11.8667 | -34.5343 | -19.746
11.8728 | -34.5347 | -19.7551
11.8789 | -34.5351 | -19.7645
11.885 | -34.5355 | -19.7734
11.8911 | -34.536 | -19.7826
11.8972 | -34.5364 | -19.7913
11.9033 | -34.5368 | -19.7999
11.9094 | -34.5372 | -19.8085
11.9155 | -34.5376 | -19.8176
11.9216 | -34.5381 | -19.8264
11.9277 | -34.5385 | -19.8352
11.9338 | -34.539 | -19.8443
11.9399 | -34.5394 | -19.8535
11.946 | -34.5398 | -19.8625
11.9521 | -34.5403 | -19.8714
11.9582 | -34.5407 | -19.8802
11.9644 | -34.5412 | -19.8892
11.9705 | -34.5416 | -19.8983
11.9766 | -34.542 | -19.9073
11.9827 | -34.5424 | -19.9161
11.9888 | -34.5428 | -19.9248
11.9949 | -34.5432 | -19.9339
12.001 | -34.5437 | -19.9428
12.0071 | -34.5441 | -19.9516
12.0132 | -34.5445 | -19.9604
12.0193 | -34.545 | -19.9695
12.0254 | -34.5454 | -19.9782
12.0315 | -34.5459 | -19.9871
12.0376 | -34.5463 | -19.9959
12.0437 | -34.5467 | -20.0047
12.0498 | -34.5471 | -20.0137
12.0559 | -34.5476 | -20.0228
12.062 | -34.548 | -20.0318
12.0681 | -34.5484 | -20.0406
12.0742 | -34.5488 | -20.0498
12.0803 | -34.5493 | -20.0586
12.0864 | -34.5497 | -20.0674
12.0925 | -34.5502 | -20.0763
12.0986 | -34.5507 | -20.0849
12.1047 | -34.5511 | -20.0938
12.1109 | -34.5515 | -20.103
12.117 | -34.552 | -20.1118
12.1231 | -34.5524 | -20.1207
12.1292 | -34.5529 | -20.1298
12.1353 | -34.5533 | -20.1385
12.1414 | -34.5538 | -20.1475
12.1475 | -34.5542 | -20.1564
12.1536 | -34.5546 | -20.1653
12.1597 | -34.555 | -20.1743
12.1658 | -34.5554 | -20.1833
12.1719 | -34.5559 | -20.1924
12.178 | -34.5564 | -20.2015
12.1841 | -34.5568 | -20.2104
12.1902 | -34.5573 | -20.2193
12.1963 | -34.5577 | -20.2281
12.2024 | -34.5582 | -20.2369
12.2085 | -34.5586 | -20.2458
12.2146 | -34.559 | -20.2546
12.2207 | -34.5594 | -20.2635
12.2268 | -34.5598 | -20.2724
12.2329 | -34.5603 | -20.2813
12.239 | -34.5606 | -20.29
12.2451 | -34.5611 | -20.2986
12.2513 | -34.5616 | -20.3075
12.2574 | -34.562 | -20.3167
12.2635 | -34.5625 | -20.3251
12.2696 | -34.5629 | -20.3339
12.2757 | -34.5633 | -20.3425
12.2818 | -34.5637 | -20.3511
12.2879 | -34.5641 | -20.3598
12.294 | -34.5645 | -20.3686
12.3001 | -34.565 | -20.3776
12.3062 | -34.5654 | -20.3865
12.3123 | -34.5659 | -20.3955
12.3184 | -34.5663 | -20.4043
12.3245 | -34.5667 | -20.4134
12.3306 | -34.5672 | -20.422
12.3367 | -34.5676 | -20.4307
12.3428 | -34.5681 | -20.4393
12.3489 | -34.5685 | -20.4478
12.355 | -34.5689 | -20.4565
12.3611 | -34.5693 | -20.4652
12.3672 | -34.5698 | -20.474
12.3733 | -34.5702 | -20.4831
12.3794 | -34.5706 | -20.4915
12.3855 | -34.571 | -20.5003
12.3916 | -34.5715 | -20.5093
12.3978 | -34.5719 | -20.5183
12.4039 | -34.5723 | -20.527
12.41 | -34.5727 | -20.536
12.4161 | -34.5732 | -20.5446
12.4222 | -34.5736 | -20.5532
12.4283 | -34.5741 | -20.5621
12.4344 | -34.5745 | -20.5708
12.4405 | -34.5749 | -20.5797
12.4466 | -34.5753 | -20.5884
12.4527 | -34.5758 | -20.5973
12.4588 | -34.5762 | -20.6061
12.4649 | -34.5767 | -20.6151
12.471 | -34.5771 | -20.6243
12.4771 | -34.5775 | -20.6326
12.4832 | -34.578 | -20.6414
12.4893 | -34.5784 | -20.65
12.4954 | -34.5789 | -20.6588
12.5015 | -34.5794 | -20.6674
12.5076 | -34.5798 | -20.6762
12.5137 | -34.5801 | -20.6853
12.5198 | -34.5806 | -20.6937
12.5259 | -34.5811 | -20.7023
12.532 | -34.5816 | -20.7114
12.5382 | -34.582 | -20.7199
12.5443 | -34.5824 | -20.7288
12.5504 | -34.5828 | -20.7376
12.5565 | -34.5833 | -20.7461
12.5626 | -34.5838 | -20.7551
12.5687 | -34.5842 | -20.7639
12.5748 | -34.5846 | -20.7726
12.5809 | -34.585 | -20.781
12.587 | -34.5854 | -20.79
12.5931 | -34.5859 | -20.7988
12.5992 | -34.5863 | -20.8077
12.6053 | -34.5867 | -20.8164
12.6114 | -34.5872 | -20.8253
12.6175 | -34.5876 | -20.8341
12.6236 | -34.5881 | -20.8426
12.6297 | -34.5885 | -20.8511
12.6358 | -34.589 | -20.8598
12.6419 | -34.5895 | -20.8688
12.648 | -34.5899 | -20.8777
12.6541 | -34.5903 | -20.8864
12.6602 | -34.5908 | -20.8948
12.6663 | -34.5912 | -20.9034
12.6724 | -34.5916 | -20.9122
12.6785 | -34.5921 | -20.9213
12.6847 | -34.5925 | -20.9302
12.6908 | -34.5929 | -20.9388
12.6969 | -34.5934 | -20.9474
12.703 | -34.5938 | -20.9563
12.7091 | -34.5943 | -20.9651
12.7152 | -34.5947 | -20.9739
12.7213 | -34.5951 | -20.9827
12.7274 | -34.5956 | -20.9916
12.7335 | -34.596 | -20.9998
12.7396 | -34.5964 | -21.0083
12.7457 | -34.5969 | -21.017
12.7518 | -34.5973 | -21.0258
12.7579 | -34.5977 | -21.0347
12.764 | -34.5981 | -21.0434
12.7701 | -34.5985 | -21.0518
12.7762 | -34.5989 | -21.0606
12.7823 | -34.5994 | -21.0693
12.7884 | -34.5997 | -21.0779
12.7945 | -34.6002 | -21.0861
12.8006 | -34.6006 | -21.0941
12.8067 | -34.6011 | -21.1031
12.8128 | -34.6015 | -21.1115
12.8189 | -34.6018 | -21.1198
12.8251 | -34.6022 | -21.1283
12.8312 | -34.6026 | -21.137
12.8373 | -34.603 | -21.1455
12.8434 | -34.6035 | -21.154
12.8495 | -34.6039 | -21.1626
12.8556 | -34.6043 | -21.1713
12.8617 | -34.6047 | -21.1797
12.8678 | -34.6052 | -21.1882
12.8739 | -34.6056 | -21.1966
12.88 | -34.6061 | -21.2053
12.8861 | -34.6064 | -21.2141
12.8922 | -34.6069 | -21.2226
12.8983 | -34.6073 | -21.2308
12.9044 | -34.6078 | -21.2394
12.9105 | -34.6083 | -21.2479
12.9166 | -34.6087 | -21.2568
12.9227 | -34.6091 | -21.2659
12.9288 | -34.6096 | -21.2745
12.9349 | -34.6101 | -21.2835
12.941 | -34.6105 | -21.2921
12.9471 | -34.611 | -21.3007
12.9532 | -34.6114 | -21.3092
12.9593 | -34.6119 | -21.3179
12.9654 | -34.6123 | -21.3265
12.9716 | -34.6127 | -21.335
12.9777 | -34.6131 | -21.3431
12.9838 | -34.6135 | -21.3516
12.9899 | -34.614 | -21.3603
12.996 | -34.6144 | -21.3691
13.0021 | -34.6148 | -21.3777
13.0082 | -34.6152 | -21.3863
13.0143 | -34.6155 | -21.3945
13.0204 | -34.6159 | -21.4029
13.0265 | -34.6164 | -21.4112
13.0326 | -34.6168 | -21.42
13.0387 | -34.6173 | -21.4286
13.0448 | -34.6177 | -21.4373
13.0509 | -34.6182 | -21.4456
13.057 | -34.6186 | -21.4539
13.0631 | -34.6191 | -21.4623
13.0692 | -34.6195 | -21.4707
13.0753 | -34.62 | -21.4791
13.0814 | -34.6204 | -21.4872
13.0875 | -34.6208 | -21.4953
13.0936 | -34.6213 | -21.5037
13.0997 | -34.6217 | -21.5119
13.1058 | -34.6221 | -21.5198
13.112 | -34.6224 | -21.5283
13.1181 | -34.6229 | -21.5367
13.1242 | -34.6233 | -21.5453
13.1303 | -34.6238 | -21.554
13.1364 | -34.6242 | -21.5625
13.1425 | -34.6246 | -21.5709
13.1486 | -34.625 | -21.5792
13.1547 | -34.6255 | -21.5872
13.1608 | -34.6259 | -21.5957
13.1669 | -34.6263 | -21.6042
13.173 | -34.6268 | -21.6126
13.1791 | -34.6272 | -21.6211
13.1852 | -34.6276 | -21.6299
13.1913 | -34.628 | -21.6382
13.1974 | -34.6284 | -21.6466
13.2035 | -34.6288 | -21.655
13.2096 | -34.6293 | -21.6634
13.2157 | -34.6297 | -21.6719
13.2218 | -34.6302 | -21.6805
13.2279 | -34.6306 | -21.6888
13.234 | -34.631 | -21.6973
13.2401 | -34.6314 | -21.7057
13.2462 | -34.6319 | -21.7142
13.2524 | -34.6323 | -21.7223
13.2585 | -34.6327 | -21.7306
13.2646 | -34.6331 | -21.7392
13.2707 | -34.6335 | -21.7477
13.2768 | -34.634 | -21.7563
13.2829 | -34.6344 | -21.7647
13.289 | -34.6348 | -21.7733
13.2951 | -34.6352 | -21.7815
13.3012 | -34.6357 | -21.79
13.3073 | -34.636 | -21.7985
13.3134 | -34.6364 | -21.8068
13.3195 | -34.6369 | -21.8151
13.3256 | -34.6373 | -21.8238
13.3317 | -34.6378 | -21.832
13.3378 | -34.6382 | -21.8405
13.3439 | -34.6386 | -21.8489
13.35 | -34.639 | -21.8572
13.3561 | -34.6395 | -21.8658
13.3622 | -34.6398 | -21.8742
13.3683 | -34.6403 | -21.8823
13.3744 | -34.6407 | -21.8906
13.3805 | -34.6411 | -21.8992
13.3866 | -34.6415 | -21.9074
13.3927 | -34.6419 | -21.9157
13.3989 | -34.6424 | -21.9239
13.405 | -34.6428 | -21.9323
13.4111 | -34.6433 | -21.9408
13.4172 | -34.6437 | -21.9489
13.4233 | -34.6441 | -21.9573
13.4294 | -34.6445 | -21.9653
13.4355 | -34.645 | -21.9743
13.4416 | -34.6454 | -21.9828
13.4477 | -34.6457 | -21.9911
13.4538 | -34.6462 | -21.9993
13.4599 | -34.6466 | -22.0078
13.466 | -34.647 | -22.0165
13.4721 | -34.6475 | -22.0248
13.4782 | -34.6479 | -22.0331
13.4843 | -34.6484 | -22.0413
13.4904 | -34.6488 | -22.0492
13.4965 | -34.6492 | -22.0579
13.5026 | -34.6496 | -22.0661
13.5087 | -34.6501 | -22.0741
13.5148 | -34.6505 | -22.0824
13.5209 | -34.6509 | -22.0905
13.527 | -34.6513 | -22.0992
13.5331 | -34.6517 | -22.1074
13.5393 | -34.6522 | -22.1158
13.5454 | -34.6525 | -22.1242
13.5515 | -34.653 | -22.1325
13.5576 | -34.6534 | -22.1409
13.5637 | -34.6539 | -22.1491
13.5698 | -34.6543 | -22.1577
13.5759 | -34.6546 | -22.1657
13.582 | -34.655 | -22.1742
13.5881 | -34.6554 | -22.1827
13.5942 | -34.6559 | -22.1909
13.6003 | -34.6563 | -22.1994
13.6064 | -34.6567 | -22.2079
13.6125 | -34.6572 | -22.2164
13.6186 | -34.6576 | -22.2247
13.6247 | -34.658 | -22.2329
13.6308 | -34.6584 | -22.2409
13.6369 | -34.6589 | -22.2492
13.643 | -34.6593 | -22.2573
13.6491 | -34.6597 | -22.2656
13.6552 | -34.6602 | -22.2737
13.6613 | -34.6606 | -22.2822
13.6674 | -34.661 | -22.2908
13.6735 | -34.6614 | -22.2991
13.6796 | -34.6618 | -22.3072
13.6858 | -34.6622 | -22.3156
13.6919 | -34.6627 | -22.3239
13.698 | -34.6631 | -22.3323
13.7041 | -34.6634 | -22.3408
13.7102 | -34.6639 | -22.3492
13.7163 | -34.6643 | -22.3575
13.7224 | -34.6648 | -22.3657
13.7285 | -34.6652 | -22.3741
13.7346 | -34.6656 | -22.3823
13.7407 | -34.666 | -22.3906
13.7468 | -34.6664 | -22.399
13.7529 | -34.6668 | -22.4072
13.759 | -34.6673 | -22.4155
13.7651 | -34.6677 | -22.4239
13.7712 | -34.6681 | -22.4324
13.7773 | -34.6685 | -22.4407
13.7834 | -34.669 | -22.4489
13.7895 | -34.6695 | -22.457
13.7956 | -34.6699 | -22.4654
13.8017 | -34.6703 | -22.4737
13.8078 | -34.6708 | -22.482
13.8139 | -34.6712 | -22.4902
13.82 | -34.6717 | -22.4983
13.8262 | -34.6721 | -22.5067
13.8323 | -34.6726 | -22.515
13.8384 | -34.673 | -22.5235
13.8445 | -34.6734 | -22.5319
13.8506 | -34.6739 | -22.5403
13.8567 | -34.6743 | -22.5487
13.8628 | -34.6747 | -22.5572
13.8689 | -34.6751 | -22.5654
13.875 | -34.6756 | -22.5737
13.8811 | -34.6761 | -22.582
13.8872 | -34.6766 | -22.5903
13.8933 | -34.677 | -22.5984
13.8994 | -34.6775 | -22.6065
13.9055 | -34.6779 | -22.6146
13.9116 | -34.6783 | -22.6226
13.9177 | -34.6788 | -22.6306
13.9238 | -34.6793 | -22.639
13.9299 | -34.6797 | -22.6472
13.936 | -34.6801 | -22.6553
13.9421 | -34.6806 | -22.6637
13.9482 | -34.6811 | -22.672
13.9543 | -34.6815 | -22.68
13.9604 | -34.682 | -22.6885
13.9665 | -34.6824 | -22.6967
13.9727 | -34.6828 | -22.7047
13.9788 | -34.6832 | -22.7132
13.9849 | -34.6836 | -22.7214
13.991 | -34.684 | -22.73
13.9971 | -34.6845 | -22.7385
14.0032 | -34.6849 | -22.7471
14.0093 | -34.6853 | -22.7551
14.0154 | -34.6857 | -22.7637
14.0215 | -34.6861 | -22.7717
14.0276 | -34.6865 | -22.7798
14.0337 | -34.687 | -22.7881
14.0398 | -34.6875 | -22.7962
14.0459 | -34.6878 | -22.8046
14.052 | -34.6882 | -22.8128
14.0581 | -34.6887 | -22.8208
14.0642 | -34.6891 | -22.8292
14.0703 | -34.6896 | -22.8374
14.0764 | -34.69 | -22.8461
14.0825 | -34.6905 | -22.8544
14.0886 | -34.6908 | -22.8622
14.0947 | -34.6912 | -22.8701
14.1008 | -34.6917 | -22.8788
14.1069 | -34.6921 | -22.8869
14.1131 | -34.6926 | -22.8953
14.1192 | -34.693 | -22.9033
14.1253 | -34.6934 | -22.9114
14.1314 | -34.6938 | -22.9196
14.1375 | -34.6943 | -22.928
14.1436 | -34.6948 | -22.9362
14.1497 | -34.6952 | -22.9442
14.1558 | -34.6956 | -22.9526
14.1619 | -34.6961 | -22.9608
14.168 | -34.6965 | -22.969
14.1741 | -34.697 | -22.977
14.1802 | -34.6974 | -22.9852
14.1863 | -34.6978 | -22.9934
14.1924 | -34.6983 | -23.0017
14.1985 | -34.6987 | -23.0098
14.2046 | -34.6991 | -23.0177
14.2107 | -34.6996 | -23.0256
14.2168 | -34.7 | -23.0342
14.2229 | -34.7005 | -23.0424
14.229 | -34.7009 | -23.0505
14.2351 | -34.7014 | -23.0585
14.2412 | -34.7018 | -23.0665
14.2473 | -34.7023 | -23.0747
14.2534 | -34.7027 | -23.0829
14.2596 | -34.7032 | -23.0912
14.2657 | -34.7036 | -23.0996
14.2718 | -34.7041 | -23.108
14.2779 | -34.7045 | -23.1158
14.284 | -34.7049 | -23.1238
14.2901 | -34.7054 | -23.1318
14.2962 | -34.7059 | -23.1397
14.3023 | -34.7063 | -23.148
14.3084 | -34.7068 | -23.1565
14.3145 | -34.7073 | -23.1646
14.3206 | -34.7077 | -23.1727
14.3267 | -34.7081 | -23.1807
14.3328 | -34.7085 | -23.1888
14.3389 | -34.709 | -23.1968
14.345 | -34.7094 | -23.2051
14.3511 | -34.7099 | -23.2129
14.3572 | -34.7103 | -23.2212
14.3633 | -34.7107 | -23.2292
14.3694 | -34.7111 | -23.2373
14.3755 | -34.7115 | -23.2451
14.3816 | -34.712 | -23.2529
14.3877 | -34.7125 | -23.2609
14.3938 | -34.7129 | -23.2691
14.4 | -34.7133 | -23.2772
14.4061 | -34.7137 | -23.2853
14.4122 | -34.7141 | -23.2937
14.4183 | -34.7146 | -23.302
14.4244 | -34.715 | -23.3098
14.4305 | -34.7154 | -23.3182
14.4366 | -34.7159 | -23.3265
14.4427 | -34.7164 | -23.3347
14.4488 | -34.7168 | -23.3425
14.4549 | -34.7172 | -23.3505
14.461 | -34.7177 | -23.3588
14.4671 | -34.7181 | -23.367
14.4732 | -34.7186 | -23.3747
14.4793 | -34.719 | -23.3834
14.4854 | -34.7195 | -23.3912
14.4915 | -34.72 | -23.3992
14.4976 | -34.7205 | -23.4074
14.5037 | -34.7209 | -23.4157
14.5098 | -34.7213 | -23.4236
14.5159 | -34.7218 | -23.4319
14.522 | -34.7222 | -23.4402
14.5281 | -34.7227 | -23.4481
14.5342 | -34.7231 | -23.4562
14.5403 | -34.7236 | -23.4645
14.5465 | -34.724 | -23.4729
14.5526 | -34.7245 | -23.4812
14.5587 | -34.7249 | -23.4893
14.5648 | -34.7254 | -23.4972
14.5709 | -34.7258 | -23.5056
14.577 | -34.7262 | -23.5138
14.5831 | -34.7267 | -23.522
14.5892 | -34.7272 | -23.5298
14.5953 | -34.7277 | -23.5376
14.6014 | -34.7282 | -23.5457
14.6075 | -34.7286 | -23.5536
14.6136 | -34.7291 | -23.5615
14.6197 | -34.7295 | -23.5698
14.6258 | -34.73 | -23.578
14.6319 | -34.7305 | -23.5858
14.638 | -34.731 | -23.5939
14.6441 | -34.7314 | -23.6015
14.6502 | -34.7318 | -23.6092
14.6563 | -34.7323 | -23.6172
14.6624 | -34.7327 | -23.6252
14.6685 | -34.7332 | -23.6329
14.6746 | -34.7336 | -23.641
14.6807 | -34.7341 | -23.6492
14.6869 | -34.7345 | -23.6572
14.693 | -34.7349 | -23.6651
14.6991 | -34.7354 | -23.6731
14.7052 | -34.7359 | -23.681
14.7113 | -34.7363 | -23.6891
14.7174 | -34.7367 | -23.697
14.7235 | -34.7372 | -23.705
14.7296 | -34.7376 | -23.713
14.7357 | -34.7381 | -23.7207
14.7418 | -34.7386 | -23.7284
14.7479 | -34.739 | -23.7366
14.754 | -34.7395 | -23.7446
14.7601 | -34.7399 | -23.7527
14.7662 | -34.7404 | -23.7609
14.7723 | -34.7408 | -23.7688
14.7784 | -34.7413 | -23.7769
14.7845 | -34.7418 | -23.785
14.7906 | -34.7422 | -23.793
14.7967 | -34.7427 | -23.8009
14.8028 | -34.7431 | -23.8089
14.8089 | -34.7436 | -23.8167
14.815 | -34.744 | -23.8245
14.8211 | -34.7444 | -23.8324
14.8272 | -34.7449 | -23.8405
14.8334 | -34.7453 | -23.8485
14.8395 | -34.7458 | -23.8565
14.8456 | -34.7462 | -23.8644
14.8517 | -34.7467 | -23.8725
14.8578 | -34.7472 | -23.8804
14.8639 | -34.7476 | -23.8886
14.87 | -34.7481 | -23.8965
14.8761 | -34.7486 | -23.9043
14.8822 | -34.749 | -23.9122
14.8883 | -34.7495 | -23.92
14.8944 | -34.75 | -23.9282
14.9005 | -34.7504 | -23.936
14.9066 | -34.7509 | -23.944
14.9127 | -34.7513 | -23.9517
14.9188 | -34.7518 | -23.96
14.9249 | -34.7523 | -23.968
14.931 | -34.7527 | -23.9758
14.9371 | -34.7532 | -23.9837
14.9432 | -34.7536 | -23.9918
14.9493 | -34.7541 | -23.9997
14.9554 | -34.7546 | -24.0076
14.9615 | -34.755 | -24.0157
14.9676 | -34.7555 | -24.0238
14.9738 | -34.756 | -24.0316
14.9799 | -34.7564 | -24.0398
14.986 | -34.757 | -24.0474
14.9921 | -34.7575 | -24.0553
14.9982 | -34.7579 | -24.0636
15.0043 | -34.7583 | -24.0715
15.0104 | -34.7588 | -24.0793
15.0165 | -34.7593 | -24.0873
15.0226 | -34.7597 | -24.095
15.0287 | -34.7602 | -24.1033
15.0348 | -34.7607 | -24.1114
15.0409 | -34.7611 | -24.1196
15.047 | -34.7616 | -24.1277
15.0531 | -34.7622 | -24.1357
15.0592 | -34.7626 | -24.1439
15.0653 | -34.7631 | -24.1518
15.0714 | -34.7636 | -24.1596
15.0775 | -34.764 | -24.1672
15.0836 | -34.7645 | -24.1753
15.0897 | -34.765 | -24.1832
15.0958 | -34.7655 | -24.1909
15.1019 | -34.7659 | -24.1989
15.108 | -34.7663 | -24.2069
15.1141 | -34.7668 | -24.2149
15.1203 | -34.7673 | -24.2228
15.1264 | -34.7678 | -24.2308
15.1325 | -34.7683 | -24.2386
15.1386 | -34.7688 | -24.2467
15.1447 | -34.7693 | -24.2546
15.1508 | -34.7698 | -24.2627
15.1569 | -34.7702 | -24.2709
15.163 | -34.7708 | -24.2785
15.1691 | -34.7713 | -24.2864
15.1752 | -34.7717 | -24.2943
15.1813 | -34.7722 | -24.3023
15.1874 | -34.7727 | -24.3103
15.1935 | -34.7733 | -24.318
15.1996 | -34.7738 | -24.3263
15.2057 | -34.7742 | -24.3341
15.2118 | -34.7747 | -24.3419
15.2179 | -34.7752 | -24.3495
15.224 | -34.7756 | -24.3576
15.2301 | -34.7761 | -24.3653
15.2362 | -34.7766 | -24.3732
15.2423 | -34.7771 | -24.3813
15.2484 | -34.7775 | -24.389
15.2545 | -34.778 | -24.3969
15.2607 | -34.7785 | -24.4048
15.2668 | -34.779 | -24.4126
15.2729 | -34.7795 | -24.4204
15.279 | -34.78 | -24.4282
15.2851 | -34.7805 | -24.4361
15.2912 | -34.7809 | -24.4438
15.2973 | -34.7814 | -24.4516
15.3034 | -34.782 | -24.4593
15.3095 | -34.7824 | -24.4669
15.3156 | -34.7828 | -24.475
15.3217 | -34.7833 | -24.4826
15.3278 | -34.7838 | -24.4905
15.3339 | -34.7843 | -24.4982
15.34 | -34.7848 | -24.506
15.3461 | -34.7852 | -24.5141
15.3522 | -34.7857 | -24.5219
15.3583 | -34.7862 | -24.5297
15.3644 | -34.7867 | -24.5377
15.3705 | -34.7872 | -24.5453
15.3766 | -34.7877 | -24.5533
15.3827 | -34.7882 | -24.561
15.3888 | -34.7886 | -24.5686
15.3949 | -34.7891 | -24.5762
15.401 | -34.7896 | -24.5842
15.4072 | -34.7901 | -24.5919
15.4133 | -34.7906 | -24.5998
15.4194 | -34.7911 | -24.6076
15.4255 | -34.7916 | -24.6154
15.4316 | -34.7921 | -24.6233
15.4377 | -34.7926 | -24.631
15.4438 | -34.7931 | -24.6383
15.4499 | -34.7936 | -24.6461
15.456 | -34.794 | -24.654
15.4621 | -34.7945 | -24.6619
15.4682 | -34.795 | -24.6698
15.4743 | -34.7956 | -24.6775
15.4804 | -34.7961 | -24.6855
15.4865 | -34.7965 | -24.6934
15.4926 | -34.797 | -24.7013
15.4987 | -34.7975 | -24.7091
15.5048 | -34.798 | -24.7168
15.5109 | -34.7985 | -24.7243
15.517 | -34.799 | -24.7322
15.5231 | -34.7995 | -24.7401
15.5292 | -34.8 | -24.7478
15.5353 | -34.8005 | -24.7556
15.5414 | -34.801 | -24.7637
15.5476 | -34.8015 | -24.7716
15.5537 | -34.802 | -24.7794
15.5598 | -34.8025 | -24.7869
15.5659 | -34.803 | -24.795
15.572 | -34.8034 | -24.8024
15.5781 | -34.8039 | -24.8103
15.5842 | -34.8045 | -24.8179
15.5903 | -34.8049 | -24.8255
15.5964 | -34.8054 | -24.8332
15.6025 | -34.8059 | -24.8409
15.6086 | -34.8064 | -24.8488
15.6147 | -34.8069 | -24.8565
15.6208 | -34.8074 | -24.8641
15.6269 | -34.808 | -24.8717
15.633 | -34.8084 | -24.8795
15.6391 | -34.8089 | -24.8874
15.6452 | -34.8094 | -24.895
15.6513 | -34.8099 | -24.903
15.6574 | -34.8104 | -24.9104
15.6635 | -34.8109 | -24.9183
15.6696 | -34.8114 | -24.9259
15.6757 | -34.8119 | -24.9341
15.6818 | -34.8124 | -24.9422
15.688 | -34.813 | -24.95
15.6941 | -34.8134 | -24.9578
15.7002 | -34.814 | -24.9655
15.7063 | -34.8144 | -24.9736
15.7124 | -34.8149 | -24.9814
15.7185 | -34.8154 | -24.9891
15.7246 | -34.8159 | -24.9966
15.7307 | -34.8164 | -25.0046
15.7368 | -34.8169 | -25.0127
15.7429 | -34.8174 | -25.0203
15.749 | -34.8179 | -25.028
15.7551 | -34.8183 | -25.0357
15.7612 | -34.8188 | -25.0428
15.7673 | -34.8193 | -25.0505
15.7734 | -34.8198 | -25.0581
15.7795 | -34.8203 | -25.0662
15.7856 | -34.8208 | -25.0738
15.7917 | -34.8213 | -25.0814
15.7978 | -34.8218 | -25.089
15.8039 | -34.8223 | -25.0967
15.81 | -34.8227 | -25.1043
15.8161 | -34.8232 | -25.1122
15.8222 | -34.8237 | -25.1199
15.8283 | -34.8242 | -25.1276
15.8345 | -34.8247 | -25.1355
15.8406 | -34.8253 | -25.1433
15.8467 | -34.8258 | -25.1509
15.8528 | -34.8263 | -25.1585
15.8589 | -34.8267 | -25.1661
15.865 | -34.8272 | -25.1738
15.8711 | -34.8277 | -25.1819
15.8772 | -34.8283 | -25.1897
15.8833 | -34.8288 | -25.1976
15.8894 | -34.8292 | -25.2053
15.8955 | -34.8297 | -25.2133
15.9016 | -34.8303 | -25.221
15.9077 | -34.8308 | -25.2288
15.9138 | -34.8313 | -25.2365
15.9199 | -34.8318 | -25.2444
15.926 | -34.8323 | -25.2525
15.9321 | -34.8328 | -25.2601
15.9382 | -34.8334 | -25.2679
15.9443 | -34.8339 | -25.2758
15.9504 | -34.8344 | -25.2834
15.9565 | -34.835 | -25.2911
15.9626 | -34.8355 | -25.2989
15.9687 | -34.836 | -25.3065
15.9749 | -34.8365 | -25.3145
15.981 | -34.837 | -25.3223
15.9871 | -34.8375 | -25.33
15.9932 | -34.838 | -25.3379
15.9993 | -34.8385 | -25.3454
16.0054 | -34.839 | -25.353
16.0115 | -34.8395 | -25.3605
16.0176 | -34.84 | -25.3684
16.0237 | -34.8405 | -25.3764
16.0298 | -34.841 | -25.3838
16.0359 | -34.8416 | -25.3914
16.042 | -34.8421 | -25.3991
16.0481 | -34.8426 | -25.4068
16.0542 | -34.8431 | -25.4147
16.0603 | -34.8436 | -25.4225
16.0664 | -34.8441 | -25.4302
16.0725 | -34.8446 | -25.4378
16.0786 | -34.8452 | -25.4459
16.0847 | -34.8457 | -25.4532
16.0908 | -34.8462 | -25.4609
16.0969 | -34.8467 | -25.4683
16.103 | -34.8472 | -25.4762
16.1091 | -34.8477 | -25.4838
16.1152 | -34.8481 | -25.4918
16.1214 | -34.8487 | -25.4996
16.1275 | -34.8492 | -25.5071
16.1336 | -34.8497 | -25.5148
16.1397 | -34.8502 | -25.5226
16.1458 | -34.8507 | -25.5303
16.1519 | -34.8512 | -25.5381
16.158 | -34.8517 | -25.546
16.1641 | -34.8523 | -25.5537
16.1702 | -34.8528 | -25.5616
16.1763 | -34.8533 | -25.5694
16.1824 | -34.8538 | -25.577
16.1885 | -34.8544 | -25.5847
16.1946 | -34.8549 | -25.5923
16.2007 | -34.8554 | -25.6003
16.2068 | -34.8559 | -25.6079
16.2129 | -34.8564 | -25.6157
16.219 | -34.8568 | -25.6233
16.2251 | -34.8573 | -25.6314
16.2312 | -34.8578 | -25.6392
16.2373 | -34.8583 | -25.647
16.2434 | -34.8588 | -25.6547
16.2495 | -34.8593 | -25.6623
16.2556 | -34.8599 | -25.67
16.2618 | -34.8603 | -25.6775
16.2679 | -34.8609 | -25.6852
16.274 | -34.8614 | -25.6931
16.2801 | -34.8619 | -25.7006
16.2862 | -34.8624 | -25.708
16.2923 | -34.8629 | -25.7158
16.2984 | -34.8634 | -25.7237
16.3045 | -34.864 | -25.7313
16.3106 | -34.8645 | -25.7386
16.3167 | -34.865 | -25.7464
16.3228 | -34.8655 | -25.7542
16.3289 | -34.866 | -25.7618
16.335 | -34.8665 | -25.7697
16.3411 | -34.867 | -25.7772
16.3472 | -34.8675 | -25.785
16.3533 | -34.868 | -25.7926
16.3594 | -34.8685 | -25.8001
16.3655 | -34.869 | -25.808
16.3716 | -34.8695 | -25.8156
16.3777 | -34.87 | -25.8233
16.3838 | -34.8705 | -25.8309
16.3899 | -34.871 | -25.8387
16.396 | -34.8715 | -25.8466
16.4021 | -34.872 | -25.8542
16.4083 | -34.8725 | -25.8619
16.4144 | -34.8731 | -25.8694
16.4205 | -34.8736 | -25.8773
16.4266 | -34.8741 | -25.8848
16.4327 | -34.8746 | -25.8924
16.4388 | -34.8751 | -25.9003
16.4449 | -34.8756 | -25.9079
16.451 | -34.8762 | -25.9157
16.4571 | -34.8767 | -25.9235
16.4632 | -34.8772 | -25.9313
16.4693 | -34.8777 | -25.939
16.4754 | -34.8782 | -25.9468
16.4815 | -34.8787 | -25.9545
16.4876 | -34.8792 | -25.9622
16.4937 | -34.8798 | -25.9699
16.4998 | -34.8803 | -25.9776
16.5059 | -34.8808 | -25.9852
16.512 | -34.8813 | -25.9931
16.5181 | -34.8818 | -26.0008
16.5242 | -34.8823 | -26.0085
16.5303 | -34.8828 | -26.0159
16.5364 | -34.8834 | -26.0235
16.5425 | -34.8839 | -26.0313
16.5487 | -34.8844 | -26.039
16.5548 | -34.8849 | -26.0466
16.5609 | -34.8854 | -26.0544
16.567 | -34.8858 | -26.062
16.5731 | -34.8864 | -26.0697
16.5792 | -34.8869 | -26.0775
16.5853 | -34.8874 | -26.085
16.5914 | -34.8879 | -26.0929
16.5975 | -34.8884 | -26.1007
16.6036 | -34.889 | -26.1082
16.6097 | -34.8895 | -26.1158
16.6158 | -34.89 | -26.1237
16.6219 | -34.8905 | -26.1312
16.628 | -34.891 | -26.139
16.6341 | -34.8915 | -26.1464
16.6402 | -34.892 | -26.1537
16.6463 | -34.8925 | -26.1614
16.6524 | -34.893 | -26.169
16.6585 | -34.8935 | -26.1768
16.6646 | -34.894 | -26.1846
16.6707 | -34.8946 | -26.192
16.6768 | -34.8951 | -26.1996
16.6829 | -34.8956 | -26.2073
16.689 | -34.8961 | -26.2151
16.6952 | -34.8966 | -26.2228
16.7013 | -34.8971 | -26.2306
16.7074 | -34.8977 | -26.2382
16.7135 | -34.8982 | -26.2459
16.7196 | -34.8987 | -26.2536
16.7257 | -34.8992 | -26.2612
16.7318 | -34.8997 | -26.2689
16.7379 | -34.9002 | -26.2768
16.744 | -34.9007 | -26.2842
16.7501 | -34.9012 | -26.292
16.7562 | -34.9017 | -26.2995
16.7623 | -34.9022 | -26.3074
16.7684 | -34.9028 | -26.3151
16.7745 | -34.9033 | -26.3225
16.7806 | -34.9038 | -26.3302
16.7867 | -34.9043 | -26.3378
16.7928 | -34.9047 | -26.3455
16.7989 | -34.9052 | -26.353
16.805 | -34.9057 | -26.3607
16.8111 | -34.9062 | -26.3683
16.8172 | -34.9067 | -26.3762
16.8233 | -34.9072 | -26.3839
16.8294 | -34.9077 | -26.3917
16.8356 | -34.9082 | -26.3995
16.8417 | -34.9088 | -26.4071
16.8478 | -34.9093 | -26.4145
16.8539 | -34.9098 | -26.4222
16.86 | -34.9103 | -26.4297
16.8661 | -34.9108 | -26.4373
16.8722 | -34.9113 | -26.4449
16.8783 | -34.9118 | -26.4523
16.8844 | -34.9123 | -26.4599
16.8905 | -34.9128 | -26.4675
16.8966 | -34.9133 | -26.4753
16.9027 | -34.9138 | -26.4827
16.9088 | -34.9144 | -26.4907
16.9149 | -34.9149 | -26.4986
16.921 | -34.9154 | -26.506
16.9271 | -34.9159 | -26.5141
16.9332 | -34.9164 | -26.5215
16.9393 | -34.9169 | -26.5294
16.9454 | -34.9174 | -26.5369
16.9515 | -34.9179 | -26.5447
16.9576 | -34.9184 | -26.5522
16.9637 | -34.919 | -26.5597
16.9698 | -34.9195 | -26.5671
16.9759 | -34.92 | -26.5744
16.9821 | -34.9205 | -26.5821
16.9882 | -34.921 | -26.59
16.9943 | -34.9215 | -26.5979
17.0004 | -34.922 | -26.6056
17.0065 | -34.9225 | -26.6132
17.0126 | -34.923 | -26.6207
17.0187 | -34.9235 | -26.6284
17.0248 | -34.924 | -26.6358
17.0309 | -34.9245 | -26.6435
17.037 | -34.925 | -26.6513
17.0431 | -34.9256 | -26.6588
17.0492 | -34.9261 | -26.6664
17.0553 | -34.9265 | -26.6742
17.0614 | -34.927 | -26.6816
17.0675 | -34.9275 | -26.6894
17.0736 | -34.928 | -26.6972
17.0797 | -34.9285 | -26.7044
17.0858 | -34.929 | -26.7121
17.0919 | -34.9295 | -26.72
17.098 | -34.93 | -26.7273
17.1041 | -34.9305 | -26.735
17.1102 | -34.931 | -26.7428
17.1163 | -34.9315 | -26.7505
17.1225 | -34.932 | -26.758
17.1286 | -34.9325 | -26.7654
17.1347 | -34.933 | -26.7729
17.1408 | -34.9335 | -26.7805
17.1469 | -34.9341 | -26.7881
17.153 | -34.9346 | -26.7957
17.1591 | -34.9351 | -26.803
17.1652 | -34.9356 | -26.8105
17.1713 | -34.9361 | -26.8181
17.1774 | -34.9366 | -26.8258
17.1835 | -34.9371 | -26.8337
17.1896 | -34.9376 | -26.8413
17.1957 | -34.938 | -26.8489
17.2018 | -34.9385 | -26.8566
17.2079 | -34.939 | -26.8641
17.214 | -34.9395 | -26.8717
17.2201 | -34.94 | -26.8793
17.2262 | -34.9405 | -26.887
17.2323 | -34.941 | -26.8946
17.2384 | -34.9416 | -26.9018
17.2445 | -34.942 | -26.9094
17.2506 | -34.9425 | -26.9171
17.2567 | -34.943 | -26.9248
17.2628 | -34.9435 | -26.9327
17.269 | -34.944 | -26.9402
17.2751 | -34.9445 | -26.9478
17.2812 | -34.9449 | -26.9554
17.2873 | -34.9454 | -26.9629
17.2934 | -34.9459 | -26.9703
17.2995 | -34.9464 | -26.9776
17.3056 | -34.9468 | -26.985
17.3117 | -34.9473 | -26.9924
17.3178 | -34.9478 | -27
17.3239 | -34.9483 | -27.0075
17.33 | -34.9488 | -27.0151
17.3361 | -34.9493 | -27.0225
17.3422 | -34.9498 | -27.03
17.3483 | -34.9503 | -27.0376
17.3544 | -34.9508 | -27.045
17.3605 | -34.9513 | -27.0526
17.3666 | -34.9518 | -27.0603
17.3727 | -34.9523 | -27.0678
17.3788 | -34.9528 | -27.0753
17.3849 | -34.9533 | -27.0828
17.391 | -34.9538 | -27.0903
17.3971 | -34.9544 | -27.0979
17.4032 | -34.9549 | -27.1057
17.4094 | -34.9553 | -27.113
17.4155 | -34.9558 | -27.1205
17.4216 | -34.9563 | -27.128
17.4277 | -34.9568 | -27.1357
17.4338 | -34.9573 | -27.1432
17.4399 | -34.9578 | -27.1506
17.446 | -34.9583 | -27.158
17.4521 | -34.9588 | -27.1654
17.4582 | -34.9593 | -27.1728
17.4643 | -34.9598 | -27.1806
17.4704 | -34.9603 | -27.188
17.4765 | -34.9607 | -27.1955
17.4826 | -34.9612 | -27.203
17.4887 | -34.9618 | -27.2107
17.4948 | -34.9623 | -27.2181
17.5009 | -34.9627 | -27.2255
17.507 | -34.9632 | -27.2331
17.5131 | -34.9637 | -27.2404
17.5192 | -34.9642 | -27.2478
17.5253 | -34.9647 | -27.2555
17.5314 | -34.9651 | -27.263
17.5375 | -34.9656 | -27.2706
17.5436 | -34.9661 | -27.2783
17.5497 | -34.9666 | -27.286
17.5559 | -34.9671 | -27.2932
17.562 | -34.9676 | -27.3007
17.5681 | -34.9681 | -27.3083
17.5742 | -34.9685 | -27.3159
17.5803 | -34.969 | -27.3232
17.5864 | -34.9695 | -27.3303
17.5925 | -34.97 | -27.338
17.5986 | -34.9705 | -27.3452
17.6047 | -34.971 | -27.3529
17.6108 | -34.9715 | -27.3601
17.6169 | -34.972 | -27.3675
17.623 | -34.9725 | -27.3749
17.6291 | -34.973 | -27.3823
17.6352 | -34.9735 | -27.3895
17.6413 | -34.9739 | -27.397
17.6474 | -34.9744 | -27.4045
17.6535 | -34.9749 | -27.4123
17.6596 | -34.9754 | -27.4196
17.6657 | -34.9759 | -27.4268
17.6718 | -34.9764 | -27.4345
17.6779 | -34.9769 | -27.4418
17.684 | -34.9774 | -27.4492
17.6901 | -34.9779 | -27.4565
17.6963 | -34.9784 | -27.4639
17.7024 | -34.9788 | -27.4712
17.7085 | -34.9793 | -27.4785
17.7146 | -34.9798 | -27.486
17.7207 | -34.9803 | -27.4934
17.7268 | -34.9808 | -27.5008
17.7329 | -34.9813 | -27.5085
17.739 | -34.9818 | -27.5163
17.7451 | -34.9823 | -27.5236
17.7512 | -34.9828 | -27.531
17.7573 | -34.9833 | -27.5385
17.7634 | -34.9838 | -27.5458
17.7695 | -34.9842 | -27.5533
17.7756 | -34.9848 | -27.5607
17.7817 | -34.9853 | -27.568
17.7878 | -34.9858 | -27.5755
17.7939 | -34.9863 | -27.5828
17.8 | -34.9868 | -27.59
17.8061 | -34.9872 | -27.5975
17.8122 | -34.9877 | -27.605
17.8183 | -34.9882 | -27.6125
17.8244 | -34.9887 | -27.62
17.8305 | -34.9892 | -27.6273
17.8366 | -34.9897 | -27.6344
17.8428 | -34.9902 | -27.6421
17.8489 | -34.9907 | -27.6493
17.855 | -34.9912 | -27.6565
17.8611 | -34.9917 | -27.6637
17.8672 | -34.9921 | -27.6709
17.8733 | -34.9926 | -27.6785
17.8794 | -34.9931 | -27.6861
17.8855 | -34.9936 | -27.6935
17.8916 | -34.9941 | -27.7006
17.8977 | -34.9946 | -27.7079
17.9038 | -34.9951 | -27.7151
17.9099 | -34.9956 | -27.7226
17.916 | -34.9961 | -27.7299
17.9221 | -34.9965 | -27.7373
17.9282 | -34.997 | -27.7448
17.9343 | -34.9975 | -27.7522
17.9404 | -34.998 | -27.7594
17.9465 | -34.9985 | -27.7669
17.9526 | -34.999 | -27.7745
17.9587 | -34.9995 | -27.7818
17.9648 | -34.9999 | -27.7891
17.9709 | -35.0004 | -27.7965
17.977 | -35.0009 | -27.8034
17.9832 | -35.0014 | -27.8108
17.9893 | -35.0019 | -27.8183
17.9954 | -35.0024 | -27.8259
18.0015 | -35.0029 | -27.8334
18.0076 | -35.0034 | -27.8402
18.0137 | -35.0039 | -27.8478
18.0198 | -35.0043 | -27.8551
18.0259 | -35.0048 | -27.8626
18.032 | -35.0053 | -27.8699
18.0381 | -35.0058 | -27.8771
18.0442 | -35.0063 | -27.8844
18.0503 | -35.0068 | -27.8919
18.0564 | -35.0072 | -27.8991
18.0625 | -35.0077 | -27.9065
18.0686 | -35.0082 | -27.9137
18.0747 | -35.0087 | -27.921
18.0808 | -35.0092 | -27.9282
18.0869 | -35.0097 | -27.9355
18.093 | -35.0102 | -27.9427
18.0991 | -35.0107 | -27.9503
18.1052 | -35.0112 | -27.9579
18.1113 | -35.0117 | -27.9649
18.1174 | -35.0122 | -27.9719
18.1236 | -35.0127 | -27.9793
18.1297 | -35.0132 | -27.9867
18.1358 | -35.0137 | -27.994
18.1419 | -35.0142 | -28.0015
18.148 | -35.0147 | -28.0089
18.1541 | -35.0152 | -28.016
18.1602 | -35.0157 | -28.0234
18.1663 | -35.0163 | -28.0307
18.1724 | -35.0168 | -28.0382
18.1785 | -35.0172 | -28.0458
18.1846 | -35.0177 | -28.0528
18.1907 | -35.0181 | -28.0603
18.1968 | -35.0186 | -28.0677
18.2029 | -35.0191 | -28.0747
18.209 | -35.0196 | -28.0822
18.2151 | -35.0201 | -28.0898
18.2212 | -35.0207 | -28.0966
18.2273 | -35.0212 | -28.1038
18.2334 | -35.0217 | -28.111
18.2395 | -35.0221 | -28.1182
18.2456 | -35.0226 | -28.1256
18.2517 | -35.0231 | -28.1329
18.2578 | -35.0235 | -28.1399
18.2639 | -35.024 | -28.1474
18.2701 | -35.0246 | -28.1547
18.2762 | -35.0251 | -28.162
18.2823 | -35.0256 | -28.1694
18.2884 | -35.0261 | -28.1768
18.2945 | -35.0266 | -28.1839
18.3006 | -35.027 | -28.1913
18.3067 | -35.0275 | -28.1985
18.3128 | -35.0281 | -28.2057
18.3189 | -35.0286 | -28.2132
18.325 | -35.0291 | -28.2202
18.3311 | -35.0295 | -28.2276
18.3372 | -35.03 | -28.2347
18.3433 | -35.0305 | -28.2418
18.3494 | -35.031 | -28.249
18.3555 | -35.0315 | -28.256
18.3616 | -35.032 | -28.2629
18.3677 | -35.0325 | -28.2698
18.3738 | -35.033 | -28.277
18.3799 | -35.0335 | -28.2841
18.386 | -35.034 | -28.2911
18.3921 | -35.0344 | -28.2982
18.3982 | -35.0349 | -28.3053
18.4043 | -35.0354 | -28.3125
18.4105 | -35.0358 | -28.3199
18.4166 | -35.0363 | -28.3272
18.4227 | -35.0368 | -28.3346
18.4288 | -35.0373 | -28.3421
18.4349 | -35.0378 | -28.3489
18.441 | -35.0383 | -28.356
18.4471 | -35.0388 | -28.3634
18.4532 | -35.0393 | -28.3708
18.4593 | -35.0398 | -28.3779
18.4654 | -35.0402 | -28.3847
18.4715 | -35.0407 | -28.3919
18.4776 | -35.0412 | -28.3992
18.4837 | -35.0417 | -28.4065
18.4898 | -35.0422 | -28.4138
18.4959 | -35.0428 | -28.4209
18.502 | -35.0433 | -28.428
18.5081 | -35.0438 | -28.4353
18.5142 | -35.0443 | -28.4429
18.5203 | -35.0447 | -28.4498
18.5264 | -35.0452 | -28.4569
18.5325 | -35.0458 | -28.4643
18.5386 | -35.0463 | -28.4715
18.5447 | -35.0468 | -28.4786
18.5508 | -35.0473 | -28.4859
18.557 | -35.0478 | -28.4933
18.5631 | -35.0482 | -28.5005
18.5692 | -35.0487 | -28.5076
18.5753 | -35.0492 | -28.5143
18.5814 | -35.0497 | -28.5217
18.5875 | -35.0502 | -28.5294
18.5936 | -35.0507 | -28.537
18.5997 | -35.0513 | -28.5443
18.6058 | -35.0517 | -28.5514
18.6119 | -35.0522 | -28.5587
18.618 | -35.0527 | -28.566
18.6241 | -35.0533 | -28.5734
18.6302 | -35.0538 | -28.5807
18.6363 | -35.0542 | -28.5877
18.6424 | -35.0547 | -28.595
18.6485 | -35.0552 | -28.6023
18.6546 | -35.0557 | -28.6093
18.6607 | -35.0562 | -28.6165
18.6668 | -35.0567 | -28.6236
18.6729 | -35.0571 | -28.6308
18.679 | -35.0576 | -28.6377
18.6851 | -35.0581 | -28.6448
18.6912 | -35.0586 | -28.6521
18.6974 | -35.0591 | -28.6593
18.7035 | -35.0596 | -28.6664
18.7096 | -35.0601 | -28.6737
18.7157 | -35.0606 | -28.6807
18.7218 | -35.0611 | -28.688
18.7279 | -35.0616 | -28.6947
18.734 | -35.0621 | -28.7017
18.7401 | -35.0626 | -28.7087
18.7462 | -35.0631 | -28.7155
18.7523 | -35.0636 | -28.7224
18.7584 | -35.0641 | -28.7295
18.7645 | -35.0646 | -28.7365
18.7706 | -35.0651 | -28.7438
18.7767 | -35.0656 | -28.7509
18.7828 | -35.0661 | -28.7579
18.7889 | -35.0666 | -28.7647
18.795 | -35.0671 | -28.7718
18.8011 | -35.0676 | -28.7789
18.8072 | -35.0681 | -28.786
18.8133 | -35.0686 | -28.7932
18.8194 | -35.0691 | -28.8004
18.8255 | -35.0696 | -28.8072
18.8316 | -35.0701 | -28.8145
18.8377 | -35.0705 | -28.8214
18.8439 | -35.071 | -28.8285
18.85 | -35.0715 | -28.8357
18.8561 | -35.072 | -28.8427
18.8622 | -35.0725 | -28.8499
18.8683 | -35.073 | -28.8567
18.8744 | -35.0735 | -28.8635
18.8805 | -35.0741 | -28.8706
18.8866 | -35.0746 | -28.8777
18.8927 | -35.0751 | -28.8846
18.8988 | -35.0756 | -28.8917
18.9049 | -35.0761 | -28.8989
18.911 | -35.0766 | -28.906
18.9171 | -35.0771 | -28.9134
18.9232 | -35.0776 | -28.9203
18.9293 | -35.078 | -28.9274
18.9354 | -35.0785 | -28.9344
18.9415 | -35.0789 | -28.9412
18.9476 | -35.0794 | -28.9487
18.9537 | -35.0799 | -28.956
18.9598 | -35.0805 | -28.9628
18.9659 | -35.081 | -28.9698
18.972 | -35.0815 | -28.9769
18.9781 | -35.082 | -28.9837
18.9843 | -35.0825 | -28.9908
18.9904 | -35.0829 | -28.9979
18.9965 | -35.0835 | -29.0048
19.0026 | -35.084 | -29.0118
19.0087 | -35.0845 | -29.0189
19.0148 | -35.085 | -29.0257
19.0209 | -35.0855 | -29.0331
19.027 | -35.0861 | -29.04
19.0331 | -35.0865 | -29.047
19.0392 | -35.0871 | -29.054
19.0453 | -35.0876 | -29.0609
19.0514 | -35.088 | -29.068
19.0575 | -35.0885 | -29.0751
19.0636 | -35.0891 | -29.082
19.0697 | -35.0896 | -29.0893
19.0758 | -35.0901 | -29.0963
19.0819 | -35.0906 | -29.1034
19.088 | -35.0911 | -29.11
19.0941 | -35.0916 | -29.1168
19.1002 | -35.0921 | -29.1239
19.1063 | -35.0926 | -29.1312
19.1124 | -35.0931 | -29.1383
19.1185 | -35.0936 | -29.1455
19.1246 | -35.0941 | -29.1524
19.1308 | -35.0946 | -29.159
19.1369 | -35.0951 | -29.1658
19.143 | -35.0956 | -29.173
19.1491 | -35.0961 | -29.1798
19.1552 | -35.0966 | -29.1871
19.1613 | -35.0971 | -29.194
19.1674 | -35.0976 | -29.2012
19.1735 | -35.0981 | -29.2082
19.1796 | -35.0986 | -29.215
19.1857 | -35.0991 | -29.2222
19.1918 | -35.0996 | -29.2292
19.1979 | -35.1 | -29.2362
19.204 | -35.1005 | -29.2429
19.2101 | -35.101 | -29.25
19.2162 | -35.1015 | -29.2573
19.2223 | -35.102 | -29.264
19.2284 | -35.1025 | -29.2713
19.2345 | -35.103 | -29.2785
19.2406 | -35.1035 | -29.2855
19.2467 | -35.104 | -29.2925
19.2528 | -35.1045 | -29.2995
19.2589 | -35.105 | -29.3066
19.265 | -35.1055 | -29.3137
19.2712 | -35.1059 | -29.3207
19.2773 | -35.1064 | -29.3275
19.2834 | -35.1069 | -29.3343
19.2895 | -35.1074 | -29.3412
19.2956 | -35.108 | -29.3484
19.3017 | -35.1085 | -29.3552
19.3078 | -35.1089 | -29.3623
19.3139 | -35.1095 | -29.3693
19.32 | -35.11 | -29.3761
19.3261 | -35.1105 | -29.3831
19.3322 | -35.111 | -29.3902
19.3383 | -35.1115 | -29.397
19.3444 | -35.1121 | -29.404
19.3505 | -35.1125 | -29.4109
19.3566 | -35.113 | -29.4178
19.3627 | -35.1135 | -29.4252
19.3688 | -35.114 | -29.4324
19.3749 | -35.1145 | -29.4394
19.381 | -35.115 | -29.4458
19.3871 | -35.1155 | -29.4527
19.3932 | -35.116 | -29.4599
19.3993 | -35.1165 | -29.4668
19.4054 | -35.117 | -29.4737
19.4115 | -35.1175 | -29.4807
19.4177 | -35.118 | -29.4876
19.4238 | -35.1185 | -29.4944
19.4299 | -35.119 | -29.5012
19.436 | -35.1196 | -29.5084
19.4421 | -35.1201 | -29.5149
19.4482 | -35.1206 | -29.5222
19.4543 | -35.1211 | -29.5288
19.4604 | -35.1215 | -29.5358
19.4665 | -35.122 | -29.5426
19.4726 | -35.1225 | -29.5494
19.4787 | -35.123 | -29.5563
19.4848 | -35.1235 | -29.5631
19.4909 | -35.124 | -29.5699
19.497 | -35.1245 | -29.5771
19.5031 | -35.125 | -29.5838
19.5092 | -35.1254 | -29.5908
19.5153 | -35.1259 | -29.598
19.5214 | -35.1264 | -29.6044
19.5275 | -35.1269 | -29.6114
19.5336 | -35.1274 | -29.6184
19.5397 | -35.1279 | -29.6253
19.5458 | -35.1284 | -29.6321
19.5519 | -35.1289 | -29.6389
19.5581 | -35.1293 | -29.6458
19.5642 | -35.1298 | -29.6529
19.5703 | -35.1304 | -29.6598
19.5764 | -35.1308 | -29.6666
19.5825 | -35.1313 | -29.6735
19.5886 | -35.1318 | -29.6804
19.5947 | -35.1323 | -29.6873
19.6008 | -35.1327 | -29.6942
19.6069 | -35.1331 | -29.7012
19.613 | -35.1337 | -29.7082
19.6191 | -35.1341 | -29.7154
19.6252 | -35.1347 | -29.7223
19.6313 | -35.1351 | -29.7287
19.6374 | -35.1356 | -29.7359
19.6435 | -35.1361 | -29.7427
19.6496 | -35.1366 | -29.7495
19.6557 | -35.1371 | -29.7562
19.6618 | -35.1376 | -29.7627
19.6679 | -35.1381 | -29.7694
19.674 | -35.1386 | -29.7761
19.6801 | -35.1391 | -29.7826
19.6862 | -35.1396 | -29.7894
19.6923 | -35.1401 | -29.7959
19.6984 | -35.1405 | -29.8024
19.7046 | -35.141 | -29.8092
19.7107 | -35.1415 | -29.8159
19.7168 | -35.142 | -29.8231
19.7229 | -35.1425 | -29.8302
19.729 | -35.1429 | -29.8368
19.7351 | -35.1434 | -29.8434
19.7412 | -35.1439 | -29.8502
19.7473 | -35.1444 | -29.8569
19.7534 | -35.1449 | -29.864
19.7595 | -35.1454 | -29.8711
19.7656 | -35.1459 | -29.8782
19.7717 | -35.1464 | -29.8848
19.7778 | -35.1469 | -29.8917
19.7839 | -35.1473 | -29.8981
19.79 | -35.1479 | -29.9048
19.7961 | -35.1484 | -29.9118
19.8022 | -35.1489 | -29.919
19.8083 | -35.1494 | -29.9256
19.8144 | -35.1499 | -29.9325
19.8205 | -35.1503 | -29.9395
19.8266 | -35.1508 | -29.9464
19.8327 | -35.1512 | -29.9536
19.8388 | -35.1517 | -29.9605
19.845 | -35.1522 | -29.9673
19.8511 | -35.1527 | -29.9743
19.8572 | -35.1531 | -29.9812
19.8633 | -35.1536 | -29.988
19.8694 | -35.1542 | -29.9951
19.8755 | -35.1547 | -30.002
19.8816 | -35.1552 | -30.0088
19.8877 | -35.1557 | -30.0158
19.8938 | -35.1562 | -30.0226
19.8999 | -35.1566 | -30.0292
19.906 | -35.1571 | -30.0363
19.9121 | -35.1576 | -30.043
19.9182 | -35.1581 | -30.0495
19.9243 | -35.1586 | -30.0567
19.9304 | -35.1591 | -30.0644
19.9365 | -35.1596 | -30.0714
19.9426 | -35.16 | -30.0786
19.9487 | -35.1605 | -30.085
19.9548 | -35.161 | -30.0923
19.9609 | -35.1616 | -30.099
19.967 | -35.162 | -30.1059
19.9731 | -35.1625 | -30.1127
19.9792 | -35.163 | -30.1194
19.9853 | -35.1636 | -30.1264
19.9915 | -35.1641 | -30.1334
19.9976 | -35.1646 | -30.1401
20.0037 | -35.1651 | -30.1465
20.0098 | -35.1656 | -30.153
20.0159 | -35.1661 | -30.1599
20.022 | -35.1666 | -30.1673
20.0281 | -35.1671 | -30.1747
20.0342 | -35.1675 | -30.1815
20.0403 | -35.168 | -30.1884
20.0464 | -35.1685 | -30.1953
20.0525 | -35.169 | -30.2022
20.0586 | -35.1695 | -30.2091
20.0647 | -35.17 | -30.2162
20.0708 | -35.1705 | -30.2231
20.0769 | -35.1709 | -30.23
20.083 | -35.1714 | -30.2368
20.0891 | -35.1719 | -30.2436
20.0952 | -35.1725 | -30.2502
20.1013 | -35.1729 | -30.257
20.1074 | -35.1734 | -30.2644
20.1135 | -35.1739 | -30.2713
20.1196 | -35.1743 | -30.2782
20.1257 | -35.1748 | -30.2851
20.1319 | -35.1752 | -30.2919
20.138 | -35.1757 | -30.2985
20.1441 | -35.1762 | -30.3051
20.1502 | -35.1767 | -30.3119
20.1563 | -35.1771 | -30.3182
20.1624 | -35.1776 | -30.3249
20.1685 | -35.1781 | -30.3318
20.1746 | -35.1786 | -30.3388
20.1807 | -35.1791 | -30.3454
20.1868 | -35.1796 | -30.3524
20.1929 | -35.18 | -30.3588
20.199 | -35.1805 | -30.3658
20.2051 | -35.181 | -30.3724
20.2112 | -35.1815 | -30.3787
20.2173 | -35.182 | -30.3856
20.2234 | -35.1825 | -30.3925
20.2295 | -35.183 | -30.399
20.2356 | -35.1835 | -30.4059
20.2417 | -35.1839 | -30.4125
20.2478 | -35.1844 | -30.4192
20.2539 | -35.1849 | -30.4261
20.26 | -35.1854 | -30.4324
20.2661 | -35.1858 | -30.4395
20.2723 | -35.1863 | -30.4462
20.2784 | -35.1868 | -30.453
20.2845 | -35.1873 | -30.4597
20.2906 | -35.1878 | -30.4664
20.2967 | -35.1883 | -30.4729
20.3028 | -35.1887 | -30.48
20.3089 | -35.1892 | -30.4867
20.315 | -35.1897 | -30.4934
20.3211 | -35.1901 | -30.5002
20.3272 | -35.1906 | -30.5064
20.3333 | -35.1911 | -30.5126
20.3394 | -35.1916 | -30.5193
20.3455 | -35.192 | -30.5261
20.3516 | -35.1925 | -30.5325
20.3577 | -35.193 | -30.5394
20.3638 | -35.1934 | -30.5464
20.3699 | -35.1939 | -30.5527
20.376 | -35.1944 | -30.5594
20.3821 | -35.1949 | -30.5659
20.3882 | -35.1953 | -30.5727
20.3943 | -35.1959 | -30.5796
20.4004 | -35.1964 | -30.5863
20.4065 | -35.1968 | -30.5929
20.4126 | -35.1973 | -30.5996
20.4188 | -35.1978 | -30.6066
20.4249 | -35.1983 | -30.6132
20.431 | -35.1988 | -30.6199
20.4371 | -35.1993 | -30.6264
20.4432 | -35.1997 | -30.6331
20.4493 | -35.2002 | -30.6398
20.4554 | -35.2007 | -30.6466
20.4615 | -35.2012 | -30.6537
20.4676 | -35.2017 | -30.6604
20.4737 | -35.2022 | -30.6662
20.4798 | -35.2027 | -30.6731
20.4859 | -35.2032 | -30.6797
20.492 | -35.2036 | -30.6866
20.4981 | -35.2041 | -30.6934
20.5042 | -35.2046 | -30.6996
20.5103 | -35.205 | -30.7065
20.5164 | -35.2055 | -30.7131
20.5225 | -35.2059 | -30.72
20.5286 | -35.2064 | -30.727
20.5347 | -35.2068 | -30.7334
20.5408 | -35.2073 | -30.7403
20.5469 | -35.2078 | -30.7466
20.553 | -35.2083 | -30.7529
20.5592 | -35.2088 | -30.7596
20.5653 | -35.2092 | -30.7664
20.5714 | -35.2096 | -30.7727
20.5775 | -35.2101 | -30.7794
20.5836 | -35.2106 | -30.7863
20.5897 | -35.2111 | -30.7931
20.5958 | -35.2116 | -30.7999
20.6019 | -35.212 | -30.8064
20.608 | -35.2125 | -30.8131
20.6141 | -35.213 | -30.8198
20.6202 | -35.2134 | -30.8266
20.6263 | -35.2138 | -30.833
20.6324 | -35.2143 | -30.8395
20.6385 | -35.2147 | -30.8462
20.6446 | -35.2152 | -30.8532
20.6507 | -35.2156 | -30.8596
20.6568 | -35.2161 | -30.8664
20.6629 | -35.2166 | -30.8728
20.669 | -35.217 | -30.8792
20.6751 | -35.2175 | -30.8863
20.6812 | -35.2179 | -30.8924
20.6873 | -35.2184 | -30.8994
20.6934 | -35.2189 | -30.9063
20.6995 | -35.2193 | -30.9129
20.7057 | -35.2197 | -30.9194
20.7118 | -35.2202 | -30.9261
20.7179 | -35.2206 | -30.9321
20.724 | -35.2211 | -30.9387
20.7301 | -35.2215 | -30.9455
20.7362 | -35.222 | -30.9526
20.7423 | -35.2225 | -30.9589
20.7484 | -35.2229 | -30.9651
20.7545 | -35.2234 | -30.9712
20.7606 | -35.2238 | -30.9777
20.7667 | -35.2243 | -30.9842
20.7728 | -35.2248 | -30.9907
20.7789 | -35.2253 | -30.9977
20.785 | -35.2258 | -31.0043
20.7911 | -35.2262 | -31.0112
20.7972 | -35.2267 | -31.018
20.8033 | -35.2271 | -31.0247
20.8094 | -35.2276 | -31.0314
20.8155 | -35.2281 | -31.0385
20.8216 | -35.2286 | -31.0449
20.8277 | -35.229 | -31.0512
20.8338 | -35.2294 | -31.0574
20.8399 | -35.2299 | -31.0638
20.8461 | -35.2304 | -31.0707
20.8522 | -35.2309 | -31.0771
20.8583 | -35.2314 | -31.0843
20.8644 | -35.2319 | -31.0908
20.8705 | -35.2324 | -31.0972
20.8766 | -35.2328 | -31.1042
20.8827 | -35.2333 | -31.1109
20.8888 | -35.2337 | -31.1173
20.8949 | -35.2341 | -31.1233
20.901 | -35.2346 | -31.1295
20.9071 | -35.235 | -31.136
20.9132 | -35.2354 | -31.1426
20.9193 | -35.2359 | -31.1493
20.9254 | -35.2364 | -31.1558
20.9315 | -35.2368 | -31.162
20.9376 | -35.2373 | -31.1688
20.9437 | -35.2378 | -31.1755
20.9498 | -35.2382 | -31.1819
20.9559 | -35.2388 | -31.1882
20.962 | -35.2393 | -31.1949
20.9681 | -35.2397 | -31.2009
20.9742 | -35.2402 | -31.2074
20.9803 | -35.2406 | -31.2136
20.9864 | -35.2411 | -31.2203
20.9926 | -35.2415 | -31.2269
20.9987 | -35.2419 | -31.2335
21.0048 | -35.2424 | -31.2401
21.0109 | -35.2428 | -31.2466
21.017 | -35.2432 | -31.253
21.0231 | -35.2436 | -31.2594
21.0292 | -35.2441 | -31.2659
21.0353 | -35.2445 | -31.2721
21.0414 | -35.245 | -31.2787
21.0475 | -35.2454 | -31.2854
21.0536 | -35.2458 | -31.2916
21.0597 | -35.2463 | -31.2978
21.0658 | -35.2467 | -31.3038
21.0719 | -35.2471 | -31.3103
21.078 | -35.2476 | -31.3171
21.0841 | -35.2481 | -31.3238
21.0902 | -35.2486 | -31.3304
21.0963 | -35.249 | -31.337
21.1024 | -35.2494 | -31.3434
21.1085 | -35.2499 | -31.3499
21.1146 | -35.2503 | -31.3563
21.1207 | -35.2507 | -31.3625
21.1268 | -35.2512 | -31.3687
21.133 | -35.2517 | -31.375
21.1391 | -35.2521 | -31.3814
21.1452 | -35.2526 | -31.388
21.1513 | -35.253 | -31.3943
21.1574 | -35.2534 | -31.4012
21.1635 | -35.2538 | -31.4073
21.1696 | -35.2543 | -31.4134
21.1757 | -35.2548 | -31.4193
21.1818 | -35.2552 | -31.4258
21.1879 | -35.2557 | -31.4325
21.194 | -35.2561 | -31.4391
21.2001 | -35.2566 | -31.4456
21.2062 | -35.257 | -31.4516
21.2123 | -35.2574 | -31.458
21.2184 | -35.2579 | -31.4643
21.2245 | -35.2583 | -31.4708
21.2306 | -35.2588 | -31.477
21.2367 | -35.2592 | -31.4833
21.2428 | -35.2597 | -31.4896
21.2489 | -35.2601 | -31.4958
21.255 | -35.2605 | -31.502
21.2611 | -35.261 | -31.5086
21.2672 | -35.2615 | -31.5149
21.2733 | -35.2619 | -31.5213
21.2795 | -35.2624 | -31.5277
21.2856 | -35.2628 | -31.5341
21.2917 | -35.2632 | -31.5405
21.2978 | -35.2636 | -31.5473
21.3039 | -35.264 | -31.5534
21.31 | -35.2645 | -31.5589
21.3161 | -35.265 | -31.5647
21.3222 | -35.2655 | -31.5709
21.3283 | -35.2659 | -31.5776
21.3344 | -35.2663 | -31.5839
21.3405 | -35.2667 | -31.5907
21.3466 | -35.2672 | -31.597
21.3527 | -35.2677 | -31.6031
21.3588 | -35.2682 | -31.6097
21.3649 | -35.2686 | -31.616
21.371 | -35.269 | -31.622
21.3771 | -35.2695 | -31.6284
21.3832 | -35.2699 | -31.6344
21.3893 | -35.2702 | -31.6405
21.3954 | -35.2707 | -31.6468
21.4015 | -35.2712 | -31.653
21.4076 | -35.2716 | -31.6593
21.4137 | -35.2721 | -31.6657
21.4199 | -35.2725 | -31.6722
21.426 | -35.2729 | -31.6785
21.4321 | -35.2734 | -31.685
21.4382 | -35.2738 | -31.6909
21.4443 | -35.2742 | -31.6977
21.4504 | -35.2746 | -31.7041
21.4565 | -35.275 | -31.7103
21.4626 | -35.2756 | -31.7168
21.4687 | -35.276 | -31.723
21.4748 | -35.2764 | -31.7294
21.4809 | -35.2769 | -31.736
21.487 | -35.2772 | -31.7423
21.4931 | -35.2777 | -31.7478
21.4992 | -35.2781 | -31.7539
21.5053 | -35.2786 | -31.7608
21.5114 | -35.279 | -31.7675
21.5175 | -35.2794 | -31.7737
21.5236 | -35.2798 | -31.7799
21.5297 | -35.2802 | -31.7855
21.5358 | -35.2807 | -31.7918
21.5419 | -35.2811 | -31.7984
21.548 | -35.2816 | -31.8044
21.5541 | -35.282 | -31.8107
21.5602 | -35.2824 | -31.817
21.5664 | -35.2829 | -31.8237
21.5725 | -35.2833 | -31.83
21.5786 | -35.2837 | -31.8359
21.5847 | -35.2841 | -31.842
21.5908 | -35.2845 | -31.8483
21.5969 | -35.285 | -31.8547
21.603 | -35.2855 | -31.8608
21.6091 | -35.286 | -31.8677
21.6152 | -35.2864 | -31.8742
21.6213 | -35.2868 | -31.8808
21.6274 | -35.2873 | -31.8872
21.6335 | -35.2878 | -31.8932
21.6396 | -35.2883 | -31.8991
21.6457 | -35.2887 | -31.9057
21.6518 | -35.2891 | -31.9116
21.6579 | -35.2895 | -31.9174
21.664 | -35.2899 | -31.9238
21.6701 | -35.2904 | -31.9303
21.6762 | -35.2908 | -31.9365
21.6823 | -35.2913 | -31.943
21.6884 | -35.2917 | -31.9492
21.6945 | -35.2921 | -31.9554
21.7006 | -35.2925 | -31.9613
21.7068 | -35.293 | -31.9682
21.7129 | -35.2934 | -31.9746
21.719 | -35.2938 | -31.9812
21.7251 | -35.2942 | -31.9879
21.7312 | -35.2947 | -31.9938
21.7373 | -35.295 | -31.9996
21.7434 | -35.2955 | -32.0058
21.7495 | -35.2959 | -32.012
21.7556 | -35.2963 | -32.0183
21.7617 | -35.2967 | -32.0247
21.7678 | -35.2971 | -32.0312
21.7739 | -35.2976 | -32.0379
21.78 | -35.298 | -32.0441
21.7861 | -35.2984 | -32.0501
21.7922 | -35.2988 | -32.0566
21.7983 | -35.2992 | -32.063
21.8044 | -35.2997 | -32.0692
21.8105 | -35.3001 | -32.0746
21.8166 | -35.3006 | -32.0809
21.8227 | -35.3009 | -32.0871
21.8288 | -35.3014 | -32.0932
21.8349 | -35.3018 | -32.0993
21.841 | -35.3023 | -32.1059
21.8471 | -35.3026 | -32.1122
21.8533 | -35.3029 | -32.1188
21.8594 | -35.3034 | -32.1256
21.8655 | -35.3039 | -32.1319
21.8716 | -35.3043 | -32.1387
21.8777 | -35.3047 | -32.1444
21.8838 | -35.3052 | -32.1501
21.8899 | -35.3056 | -32.1565
21.896 | -35.306 | -32.163
21.9021 | -35.3064 | -32.1691
21.9082 | -35.3068 | -32.1754
21.9143 | -35.3073 | -32.1814
21.9204 | -35.3076 | -32.1876
21.9265 | -35.3081 | -32.1938
21.9326 | -35.3085 | -32.2002
21.9387 | -35.309 | -32.2065
21.9448 | -35.3094 | -32.2126
21.9509 | -35.3098 | -32.2182
21.957 | -35.3102 | -32.2243
21.9631 | -35.3107 | -32.2308
21.9692 | -35.3111 | -32.2375
21.9753 | -35.3115 | -32.2437
21.9814 | -35.3119 | -32.2498
21.9875 | -35.3123 | -32.2555
21.9937 | -35.3128 | -32.262
21.9998 | -35.3132 | -32.2682
22.0059 | -35.3137 | -32.2744
22.012 | -35.3141 | -32.2806
22.0181 | -35.3146 | -32.2862
22.0242 | -35.315 | -32.2923
22.0303 | -35.3155 | -32.2988
22.0364 | -35.3159 | -32.305
22.0425 | -35.3162 | -32.3112
22.0486 | -35.3167 | -32.317
22.0547 | -35.3171 | -32.3231
22.0608 | -35.3175 | -32.3292
22.0669 | -35.3179 | -32.3349
22.073 | -35.3183 | -32.3411
22.0791 | -35.3187 | -32.3473
22.0852 | -35.3191 | -32.3537
22.0913 | -35.3195 | -32.3596
22.0974 | -35.3199 | -32.3656
22.1035 | -35.3203 | -32.3718
22.1096 | -35.3207 | -32.3781
22.1157 | -35.3212 | -32.384
22.1218 | -35.3216 | -32.3902
22.1279 | -35.3221 | -32.3962
22.134 | -35.3226 | -32.4026
22.1402 | -35.3229 | -32.4088
22.1463 | -35.3234 | -32.4147
22.1524 | -35.3238 | -32.4209
22.1585 | -35.3242 | -32.4272
22.1646 | -35.3246 | -32.4331
22.1707 | -35.3249 | -32.4393
22.1768 | -35.3254 | -32.4453
22.1829 | -35.3258 | -32.4513
22.189 | -35.3262 | -32.457
22.1951 | -35.3266 | -32.4629
22.2012 | -35.327 | -32.4691
22.2073 | -35.3274 | -32.4751
22.2134 | -35.3278 | -32.4813
22.2195 | -35.3283 | -32.4877
22.2256 | -35.3287 | -32.4939
22.2317 | -35.3291 | -32.5
22.2378 | -35.3295 | -32.5059
22.2439 | -35.33 | -32.5124
22.25 | -35.3304 | -32.5188
22.2561 | -35.3308 | -32.5252
22.2622 | -35.3312 | -32.5317
22.2683 | -35.3316 | -32.5376
22.2744 | -35.332 | -32.5438
22.2806 | -35.3324 | -32.5497
22.2867 | -35.3328 | -32.5556
22.2928 | -35.3332 | -32.562
22.2989 | -35.3336 | -32.5683
22.305 | -35.3339 | -32.5736
22.3111 | -35.3343 | -32.5801
22.3172 | -35.3348 | -32.5862
22.3233 | -35.3352 | -32.592
22.3294 | -35.3356 | -32.5983
22.3355 | -35.336 | -32.6042
22.3416 | -35.3364 | -32.6102
22.3477 | -35.3368 | -32.6166
22.3538 | -35.3373 | -32.6229
22.3599 | -35.3377 | -32.6286
22.366 | -35.3381 | -32.6349
22.3721 | -35.3385 | -32.6404
22.3782 | -35.3389 | -32.6467
22.3843 | -35.3394 | -32.6529
22.3904 | -35.3398 | -32.6593
22.3965 | -35.3402 | -32.6653
22.4026 | -35.3407 | -32.6715
22.4087 | -35.341 | -32.6774
22.4148 | -35.3414 | -32.6835
22.4209 | -35.3418 | -32.6898
22.4271 | -35.3422 | -32.6962
22.4332 | -35.3426 | -32.7017
22.4393 | -35.343 | -32.7076
22.4454 | -35.3434 | -32.7137
22.4515 | -35.3438 | -32.7201
22.4576 | -35.3442 | -32.7262
22.4637 | -35.3447 | -32.7325
22.4698 | -35.3451 | -32.7386
22.4759 | -35.3454 | -32.7444
22.482 | -35.3459 | -32.7504
22.4881 | -35.3462 | -32.7565
22.4942 | -35.3467 | -32.7624
22.5003 | -35.3471 | -32.7685
22.5064 | -35.3474 | -32.7746
22.5125 | -35.3478 | -32.7806
22.5186 | -35.3482 | -32.7866
22.5247 | -35.3486 | -32.7926
22.5308 | -35.3491 | -32.799
22.5369 | -35.3495 | -32.8049
22.543 | -35.3499 | -32.8112
22.5491 | -35.3503 | -32.817
22.5552 | -35.3507 | -32.823
22.5613 | -35.3511 | -32.8293
22.5675 | -35.3515 | -32.8357
22.5736 | -35.3519 | -32.842
22.5797 | -35.3523 | -32.8478
22.5858 | -35.3527 | -32.8534
22.5919 | -35.353 | -32.8593
22.598 | -35.3535 | -32.8654
22.6041 | -35.3539 | -32.8714
22.6102 | -35.3543 | -32.8776
22.6163 | -35.3547 | -32.8835
22.6224 | -35.355 | -32.8897
22.6285 | -35.3555 | -32.8957
22.6346 | -35.3559 | -32.9019
22.6407 | -35.3562 | -32.9081
22.6468 | -35.3566 | -32.9143
22.6529 | -35.3569 | -32.9207
22.659 | -35.3574 | -32.9265
22.6651 | -35.3578 | -32.9321
22.6712 | -35.3582 | -32.938
22.6773 | -35.3586 | -32.9435
22.6834 | -35.3591 | -32.9497
22.6895 | -35.3595 | -32.9557
22.6956 | -35.3599 | -32.9619
22.7017 | -35.3603 | -32.9678
22.7079 | -35.3607 | -32.9743
22.714 | -35.3612 | -32.9803
22.7201 | -35.3616 | -32.9864
22.7262 | -35.362 | -32.9924
22.7323 | -35.3624 | -32.998
22.7384 | -35.3629 | -33.0039
22.7445 | -35.3633 | -33.0096
22.7506 | -35.3637 | -33.0153
22.7567 | -35.3641 | -33.0214
22.7628 | -35.3645 | -33.0273
22.7689 | -35.3649 | -33.0333
22.775 | -35.3652 | -33.0395
22.7811 | -35.3656 | -33.0455
22.7872 | -35.366 | -33.0515
22.7933 | -35.3663 | -33.0573
22.7994 | -35.3668 | -33.0631
22.8055 | -35.3672 | -33.0689
22.8116 | -35.3676 | -33.075
22.8177 | -35.368 | -33.0811
22.8238 | -35.3684 | -33.0873
22.8299 | -35.3687 | -33.0929
22.836 | -35.3692 | -33.0989
22.8421 | -35.3696 | -33.1047
22.8482 | -35.3699 | -33.1107
22.8544 | -35.3703 | -33.1167
22.8605 | -35.3708 | -33.1227
22.8666 | -35.3711 | -33.1287
22.8727 | -35.3715 | -33.1348
22.8788 | -35.3719 | -33.1408
22.8849 | -35.3723 | -33.1467
22.891 | -35.3727 | -33.1527
22.8971 | -35.3732 | -33.1586
22.9032 | -35.3736 | -33.1648
22.9093 | -35.374 | -33.1706
22.9154 | -35.3743 | -33.1768
22.9215 | -35.3747 | -33.1828
22.9276 | -35.3751 | -33.1888
22.9337 | -35.3755 | -33.1944
22.9398 | -35.3759 | -33.2004
22.9459 | -35.3763 | -33.2062
22.952 | -35.3767 | -33.2122
22.9581 | -35.377 | -33.2181
22.9642 | -35.3774 | -33.2242
22.9703 | -35.3778 | -33.2299
22.9764 | -35.3782 | -33.236
22.9825 | -35.3786 | -33.242
22.9886 | -35.3791 | -33.248
22.9948 | -35.3795 | -33.2541
23.0009 | -35.3798 | -33.26
23.007 | -35.3803 | -33.2659
23.0131 | -35.3807 | -33.2719
23.0192 | -35.3811 | -33.2777
23.0253 | -35.3815 | -33.2839
23.0314 | -35.382 | -33.2898
23.0375 | -35.3824 | -33.2958
23.0436 | -35.3827 | -33.3017
23.0497 | -35.3832 | -33.3077
23.0558 | -35.3836 | -33.3135
23.0619 | -35.384 | -33.3197
23.068 | -35.3844 | -33.3258
23.0741 | -35.3848 | -33.3315
23.0802 | -35.3852 | -33.3372
23.0863 | -35.3856 | -33.343
23.0924 | -35.3859 | -33.3489
23.0985 | -35.3862 | -33.3549
23.1046 | -35.3867 | -33.3606
23.1107 | -35.387 | -33.3666
23.1168 | -35.3874 | -33.3726
23.1229 | -35.3878 | -33.3787
23.129 | -35.3882 | -33.3849
23.1351 | -35.3886 | -33.3911
23.1413 | -35.389 | -33.3972
23.1474 | -35.3894 | -33.403
23.1535 | -35.3898 | -33.4089
23.1596 | -35.3902 | -33.4146
23.1657 | -35.3905 | -33.4202
23.1718 | -35.3909 | -33.4262
23.1779 | -35.3913 | -33.4322
23.184 | -35.3917 | -33.4379
23.1901 | -35.3921 | -33.4438
23.1962 | -35.3925 | -33.4497
23.2023 | -35.3929 | -33.4555
23.2084 | -35.3934 | -33.4614
23.2145 | -35.3937 | -33.4674
23.2206 | -35.3942 | -33.4737
23.2267 | -35.3946 | -33.4799
23.2328 | -35.395 | -33.486
23.2389 | -35.3953 | -33.4922
23.245 | -35.3957 | -33.4983
23.2511 | -35.3961 | -33.5043
23.2572 | -35.3965 | -33.5098
23.2633 | -35.3969 | -33.5159
23.2694 | -35.3973 | -33.5219
23.2755 | -35.3977 | -33.528
23.2817 | -35.3981 | -33.534
23.2878 | -35.3985 | -33.5399
23.2939 | -35.3989 | -33.5458
23.3 | -35.3993 | -33.5518
23.3061 | -35.3997 | -33.5579
23.3122 | -35.4001 | -33.5637
23.3183 | -35.4005 | -33.5694
23.3244 | -35.4009 | -33.5755
23.3305 | -35.4013 | -33.5818
23.3366 | -35.4017 | -33.5877
23.3427 | -35.402 | -33.5935
23.3488 | -35.4025 | -33.5992
23.3549 | -35.4028 | -33.6052
23.361 | -35.4033 | -33.611
23.3671 | -35.4037 | -33.6172
23.3732 | -35.4041 | -33.6231
23.3793 | -35.4045 | -33.6288
23.3854 | -35.4049 | -33.6346
23.3915 | -35.4053 | -33.6405
23.3976 | -35.4056 | -33.6464
23.4037 | -35.406 | -33.6524
23.4098 | -35.4064 | -33.6581
23.4159 | -35.4068 | -33.6638
23.422 | -35.4072 | -33.6696
23.4282 | -35.4076 | -33.6754
23.4343 | -35.408 | -33.6811
23.4404 | -35.4084 | -33.6873
23.4465 | -35.4088 | -33.693
23.4526 | -35.4092 | -33.6991
23.4587 | -35.4096 | -33.7049
23.4648 | -35.41 | -33.7107
23.4709 | -35.4103 | -33.7167
23.477 | -35.4107 | -33.7228
23.4831 | -35.4111 | -33.7288
23.4892 | -35.4115 | -33.7349
23.4953 | -35.4119 | -33.7409
23.5014 | -35.4123 | -33.7467
23.5075 | -35.4127 | -33.7526
23.5136 | -35.4131 | -33.7582
23.5197 | -35.4135 | -33.7643
23.5258 | -35.4139 | -33.7704
23.5319 | -35.4144 | -33.776
23.538 | -35.4147 | -33.7819
23.5441 | -35.4151 | -33.7881
23.5502 | -35.4155 | -33.7943
23.5563 | -35.4159 | -33.8003
23.5624 | -35.4163 | -33.8063
23.5686 | -35.4168 | -33.8124
23.5747 | -35.4172 | -33.8182
23.5808 | -35.4176 | -33.8243
23.5869 | -35.418 | -33.8301
23.593 | -35.4184 | -33.836
23.5991 | -35.4188 | -33.8419
23.6052 | -35.4192 | -33.8474
23.6113 | -35.4196 | -33.8532
23.6174 | -35.42 | -33.8591
23.6235 | -35.4204 | -33.865
23.6296 | -35.4208 | -33.8713
23.6357 | -35.4212 | -33.8774
23.6418 | -35.4216 | -33.883
23.6479 | -35.422 | -33.889
23.654 | -35.4224 | -33.895
23.6601 | -35.4228 | -33.9008
23.6662 | -35.4233 | -33.9067
23.6723 | -35.4236 | -33.9125
23.6784 | -35.4241 | -33.9184
23.6845 | -35.4244 | -33.9245
23.6906 | -35.4248 | -33.9305
23.6967 | -35.4252 | -33.9362
23.7028 | -35.4256 | -33.9422
23.7089 | -35.426 | -33.9484
23.7151 | -35.4264 | -33.9543
23.7212 | -35.4268 | -33.9601
23.7273 | -35.4272 | -33.9659
23.7334 | -35.4276 | -33.9716
23.7395 | -35.428 | -33.9775
23.7456 | -35.4284 | -33.9838
23.7517 | -35.4288 | -33.9898
23.7578 | -35.4292 | -33.9957
23.7639 | -35.4296 | -34.0017
23.77 | -35.43 | -34.0074
23.7761 | -35.4304 | -34.0131
23.7822 | -35.4308 | -34.0193
23.7883 | -35.4312 | -34.0251
23.7944 | -35.4316 | -34.0313
23.8005 | -35.432 | -34.0372
23.8066 | -35.4324 | -34.0432
23.8127 | -35.4328 | -34.049
23.8188 | -35.4332 | -34.0549
23.8249 | -35.4336 | -34.0608
23.831 | -35.434 | -34.0668
23.8371 | -35.4344 | -34.0723
23.8432 | -35.4348 | -34.0783
23.8493 | -35.4352 | -34.0837
23.8555 | -35.4357 | -34.0898
23.8616 | -35.4361 | -34.0957
23.8677 | -35.4365 | -34.1016
23.8738 | -35.4369 | -34.1075
23.8799 | -35.4373 | -34.1134
23.886 | -35.4377 | -34.1191
23.8921 | -35.4381 | -34.1253
23.8982 | -35.4385 | -34.1313
23.9043 | -35.4389 | -34.1374
23.9104 | -35.4393 | -34.1433
23.9165 | -35.4397 | -34.149
23.9226 | -35.4401 | -34.1549
23.9287 | -35.4405 | -34.1607
23.9348 | -35.4409 | -34.1666
23.9409 | -35.4413 | -34.1726
23.947 | -35.4416 | -34.1786
23.9531 | -35.4421 | -34.1841
23.9592 | -35.4425 | -34.1898
23.9653 | -35.4429 | -34.1958
23.9714 | -35.4433 | -34.2016
23.9775 | -35.4438 | -34.2074
23.9836 | -35.4442 | -34.2133
23.9897 | -35.4446 | -34.2188
23.9958 | -35.445 | -34.2247
24.002 | -35.4454 | -34.2305
24.0081 | -35.4458 | -34.2365
24.0142 | -35.4462 | -34.2425
24.0203 | -35.4466 | -34.2482
24.0264 | -35.447 | -34.2543
24.0325 | -35.4474 | -34.2602
24.0386 | -35.4478 | -34.266
24.0447 | -35.4481 | -34.2716
24.0508 | -35.4485 | -34.2773
24.0569 | -35.449 | -34.2834
24.063 | -35.4494 | -34.2892
24.0691 | -35.4497 | -34.2948
24.0752 | -35.4502 | -34.3007
24.0813 | -35.4506 | -34.3067
24.0874 | -35.451 | -34.3127
24.0935 | -35.4514 | -34.3185
24.0996 | -35.4517 | -34.3244
24.1057 | -35.4521 | -34.3304
24.1118 | -35.4525 | -34.3362
24.1179 | -35.4529 | -34.3425
24.124 | -35.4533 | -34.3484
24.1301 | -35.4538 | -34.3543
24.1362 | -35.4543 | -34.3599
24.1424 | -35.4547 | -34.3658
24.1485 | -35.4552 | -34.3712
24.1546 | -35.4555 | -34.3774
24.1607 | -35.456 | -34.3834
24.1668 | -35.4564 | -34.3892
24.1729 | -35.4569 | -34.3952
24.179 | -35.4572 | -34.4011
24.1851 | -35.4575 | -34.4068
24.1912 | -35.4579 | -34.4131
24.1973 | -35.4583 | -34.419
24.2034 | -35.4588 | -34.4244
24.2095 | -35.4593 | -34.4301
24.2156 | -35.4597 | -34.4362
24.2217 | -35.46 | -34.4425
24.2278 | -35.4604 | -34.4483
24.2339 | -35.4608 | -34.454
24.24 | -35.4612 | -34.4598
24.2461 | -35.4616 | -34.4653
24.2522 | -35.462 | -34.4709
24.2583 | -35.4624 | -34.477
24.2644 | -35.4628 | -34.4831
24.2705 | -35.4633 | -34.489
24.2766 | -35.4637 | -34.4947
24.2827 | -35.4641 | -34.5002
24.2889 | -35.4646 | -34.5059
24.295 | -35.465 | -34.5119
24.3011 | -35.4654 | -34.5178
24.3072 | -35.4658 | -34.5237
24.3133 | -35.4662 | -34.5295
24.3194 | -35.4666 | -34.5356
24.3255 | -35.467 | -34.5416
24.3316 | -35.4674 | -34.5475
24.3377 | -35.4678 | -34.5533
24.3438 | -35.4681 | -34.5594
24.3499 | -35.4685 | -34.5656
24.356 | -35.4689 | -34.5714
24.3621 | -35.4693 | -34.5771
24.3682 | -35.4697 | -34.5827
24.3743 | -35.47 | -34.5888
24.3804 | -35.4704 | -34.5944
24.3865 | -35.4708 | -34.6005
24.3926 | -35.4712 | -34.6063
24.3987 | -35.4716 | -34.6124
24.4048 | -35.472 | -34.6183
24.4109 | -35.4724 | -34.6243
24.417 | -35.4728 | -34.6299
24.4231 | -35.4732 | -34.6357
24.4293 | -35.4736 | -34.6414
24.4354 | -35.474 | -34.6473
24.4415 | -35.4744 | -34.653
24.4476 | -35.4749 | -34.6587
24.4537 | -35.4753 | -34.6643
24.4598 | -35.4758 | -34.6701
24.4659 | -35.4761 | -34.676
24.472 | -35.4765 | -34.6818
24.4781 | -35.477 | -34.6879
24.4842 | -35.4774 | -34.6936
24.4903 | -35.4778 | -34.6994
24.4964 | -35.4782 | -34.7052
24.5025 | -35.4785 | -34.7112
24.5086 | -35.4789 | -34.7169
24.5147 | -35.4794 | -34.7223
24.5208 | -35.4798 | -34.7279
24.5269 | -35.4802 | -34.734
24.533 | -35.4807 | -34.7399
24.5391 | -35.4811 | -34.7455
24.5452 | -35.4815 | -34.7514
24.5513 | -35.482 | -34.757
24.5574 | -35.4824 | -34.7629
24.5635 | -35.4828 | -34.7686
24.5696 | -35.4832 | -34.7745
24.5758 | -35.4837 | -34.7803
24.5819 | -35.4841 | -34.7861
24.588 | -35.4845 | -34.7918
24.5941 | -35.4849 | -34.7972
24.6002 | -35.4853 | -34.8026
24.6063 | -35.4857 | -34.8085
24.6124 | -35.4861 | -34.8145
24.6185 | -35.4865 | -34.8203
24.6246 | -35.4869 | -34.8257
24.6307 | -35.4873 | -34.8315
24.6368 | -35.4877 | -34.8375
24.6429 | -35.4881 | -34.8432
24.649 | -35.4885 | -34.8489
24.6551 | -35.4889 | -34.8547
24.6612 | -35.4894 | -34.8605
24.6673 | -35.4898 | -34.866
24.6734 | -35.4902 | -34.8717
24.6795 | -35.4906 | -34.8775
24.6856 | -35.491 | -34.8834
24.6917 | -35.4914 | -34.8888
24.6978 | -35.4919 | -34.8943
24.7039 | -35.4922 | -34.8998
24.71 | -35.4927 | -34.9057
24.7162 | -35.4931 | -34.9114
24.7223 | -35.4935 | -34.9173
24.7284 | -35.4939 | -34.9226
24.7345 | -35.4943 | -34.9286
24.7406 | -35.4947 | -34.9345
24.7467 | -35.4952 | -34.9402
24.7528 | -35.4955 | -34.9457
24.7589 | -35.4959 | -34.9514
24.765 | -35.4963 | -34.9568
24.7711 | -35.4967 | -34.9624
24.7772 | -35.4971 | -34.9683
24.7833 | -35.4975 | -34.9742
24.7894 | -35.4979 | -34.9798
24.7955 | -35.4983 | -34.9856
24.8016 | -35.4987 | -34.9911
24.8077 | -35.4991 | -34.9974
24.8138 | -35.4995 | -35.003
24.8199 | -35.4999 | -35.0087
24.826 | -35.5003 | -35.0144
24.8321 | -35.5007 | -35.02
24.8382 | -35.5011 | -35.0256
24.8443 | -35.5015 | -35.0313
24.8504 | -35.5019 | -35.037
24.8565 | -35.5023 | -35.0428
24.8627 | -35.5027 | -35.0486
24.8688 | -35.5032 | -35.0545
24.8749 | -35.5036 | -35.0601
24.881 | -35.504 | -35.066
24.8871 | -35.5045 | -35.072
24.8932 | -35.5049 | -35.0775
24.8993 | -35.5053 | -35.0831
24.9054 | -35.5057 | -35.089
24.9115 | -35.5061 | -35.0946
24.9176 | -35.5065 | -35.1004
24.9237 | -35.5069 | -35.1062
24.9298 | -35.5073 | -35.1118
24.9359 | -35.5077 | -35.1178
24.942 | -35.5082 | -35.1239
24.9481 | -35.5086 | -35.1299
24.9542 | -35.509 | -35.1355
24.9603 | -35.5095 | -35.141
24.9664 | -35.5098 | -35.1468
24.9725 | -35.5103 | -35.1525
24.9786 | -35.5107 | -35.1579
24.9847 | -35.5111 | -35.1635
24.9908 | -35.5115 | -35.1694
24.9969 | -35.5119 | -35.1752
25.0031 | -35.5123 | -35.1808
25.0092 | -35.5128 | -35.1864
25.0153 | -35.5132 | -35.1921
25.0214 | -35.5136 | -35.1982
25.0275 | -35.514 | -35.2039
25.0336 | -35.5144 | -35.2094
25.0397 | -35.5148 | -35.2151
25.0458 | -35.5152 | -35.2205
25.0519 | -35.5156 | -35.2263
25.058 | -35.516 | -35.2322
25.0641 | -35.5164 | -35.2378
25.0702 | -35.5168 | -35.2437
25.0763 | -35.5171 | -35.2492
25.0824 | -35.5176 | -35.2548
25.0885 | -35.518 | -35.2607
25.0946 | -35.5184 | -35.2663
25.1007 | -35.5187 | -35.2719
25.1068 | -35.5192 | -35.2774
25.1129 | -35.5195 | -35.2831
25.119 | -35.5199 | -35.2885
25.1251 | -35.5203 | -35.2942
25.1312 | -35.5207 | -35.3
25.1373 | -35.5212 | -35.3059
25.1435 | -35.5216 | -35.3115
25.1496 | -35.522 | -35.317
25.1557 | -35.5224 | -35.3229
25.1618 | -35.5229 | -35.3288
25.1679 | -35.5233 | -35.3344
25.174 | -35.5237 | -35.34
25.1801 | -35.5241 | -35.3455
25.1862 | -35.5246 | -35.3512
25.1923 | -35.525 | -35.3569
25.1984 | -35.5254 | -35.3625
25.2045 | -35.5258 | -35.3679
25.2106 | -35.5262 | -35.3738
25.2167 | -35.5266 | -35.3797
25.2228 | -35.527 | -35.3852
25.2289 | -35.5274 | -35.3906
25.235 | -35.5278 | -35.3963
25.2411 | -35.5282 | -35.4019
25.2472 | -35.5286 | -35.4075
25.2533 | -35.529 | -35.4132
25.2594 | -35.5294 | -35.4188
25.2655 | -35.5298 | -35.4244
25.2716 | -35.5302 | -35.4302
25.2777 | -35.5307 | -35.4361
25.2838 | -35.531 | -35.4419
25.29 | -35.5314 | -35.4474
25.2961 | -35.5318 | -35.4534
25.3022 | -35.5322 | -35.459
25.3083 | -35.5326 | -35.4644
25.3144 | -35.533 | -35.4701
25.3205 | -35.5334 | -35.4757
25.3266 | -35.5338 | -35.4813
25.3327 | -35.5342 | -35.4871
25.3388 | -35.5346 | -35.4927
25.3449 | -35.535 | -35.4986
25.351 | -35.5354 | -35.5043
25.3571 | -35.5358 | -35.5101
25.3632 | -35.5362 | -35.5154
25.3693 | -35.5366 | -35.5211
25.3754 | -35.537 | -35.5271
25.3815 | -35.5374 | -35.5328
25.3876 | -35.5379 | -35.5386
25.3937 | -35.5383 | -35.5443
25.3998 | -35.5387 | -35.5496
25.4059 | -35.5391 | -35.5554
25.412 | -35.5395 | -35.5609
25.4181 | -35.5399 | -35.5667
25.4242 | -35.5403 | -35.5723
25.4304 | -35.5407 | -35.5778
25.4365 | -35.5412 | -35.5833
25.4426 | -35.5416 | -35.5889
25.4487 | -35.542 | -35.5944
25.4548 | -35.5424 | -35.6002
25.4609 | -35.5428 | -35.6058
25.467 | -35.5433 | -35.6116
25.4731 | -35.5437 | -35.6173
25.4792 | -35.5441 | -35.6229
25.4853 | -35.5445 | -35.6283
25.4914 | -35.5449 | -35.6338
25.4975 | -35.5453 | -35.6393
25.5036 | -35.5457 | -35.6446
25.5097 | -35.5461 | -35.6502
25.5158 | -35.5465 | -35.656
25.5219 | -35.547 | -35.6614
25.528 | -35.5474 | -35.6668
25.5341 | -35.5478 | -35.6725
25.5402 | -35.5482 | -35.6777
25.5463 | -35.5486 | -35.6835
25.5524 | -35.549 | -35.6895
25.5585 | -35.5494 | -35.6949
25.5646 | -35.5498 | -35.7007
25.5707 | -35.5502 | -35.7061
25.5769 | -35.5505 | -35.7114
25.583 | -35.5509 | -35.7171
25.5891 | -35.5513 | -35.7226
25.5952 | -35.5517 | -35.728
25.6013 | -35.5521 | -35.7336
25.6074 | -35.5525 | -35.7389
25.6135 | -35.553 | -35.7445
25.6196 | -35.5534 | -35.7502
25.6257 | -35.5538 | -35.7561
25.6318 | -35.5542 | -35.7617
25.6379 | -35.5546 | -35.7673
25.644 | -35.555 | -35.7721
25.6501 | -35.5554 | -35.7777
25.6562 | -35.5558 | -35.7832
25.6623 | -35.5562 | -35.7888
25.6684 | -35.5566 | -35.7943
25.6745 | -35.557 | -35.8
25.6806 | -35.5574 | -35.8055
25.6867 | -35.5578 | -35.8112
25.6928 | -35.5582 | -35.8168
25.6989 | -35.5586 | -35.8224
25.705 | -35.559 | -35.8281
25.7111 | -35.5594 | -35.8336
25.7173 | -35.5598 | -35.839
25.7234 | -35.5602 | -35.8444
25.7295 | -35.5607 | -35.8501
25.7356 | -35.5611 | -35.8557
25.7417 | -35.5615 | -35.8612
25.7478 | -35.5619 | -35.8667
25.7539 | -35.5623 | -35.8723
25.76 | -35.5627 | -35.8782
25.7661 | -35.5631 | -35.884
25.7722 | -35.5635 | -35.8897
25.7783 | -35.5639 | -35.8954
25.7844 | -35.5643 | -35.9007
25.7905 | -35.5648 | -35.9063
25.7966 | -35.5652 | -35.9118
25.8027 | -35.5656 | -35.9176
25.8088 | -35.566 | -35.9233
25.8149 | -35.5664 | -35.9288
25.821 | -35.5668 | -35.9344
25.8271 | -35.5672 | -35.9401
25.8332 | -35.5676 | -35.9457
25.8393 | -35.568 | -35.9513
25.8454 | -35.5684 | -35.957
25.8515 | -35.5688 | -35.9627
25.8576 | -35.5692 | -35.968
25.8638 | -35.5696 | -35.9735
25.8699 | -35.5699 | -35.9794
25.876 | -35.5703 | -35.9851
25.8821 | -35.5708 | -35.9908
25.8882 | -35.5712 | -35.9962
25.8943 | -35.5716 | -36.002
25.9004 | -35.572 | -36.0075
25.9065 | -35.5724 | -36.0129
25.9126 | -35.5728 | -36.0182
25.9187 | -35.5732 | -36.0238
25.9248 | -35.5736 | -36.0294
25.9309 | -35.574 | -36.035
25.937 | -35.5744 | -36.0408
25.9431 | -35.5749 | -36.0463
25.9492 | -35.5753 | -36.0518
25.9553 | -35.5757 | -36.0572
25.9614 | -35.5761 | -36.0626
25.9675 | -35.5765 | -36.0679
25.9736 | -35.5769 | -36.0731
25.9797 | -35.5773 | -36.0788
25.9858 | -35.5777 | -36.0842
25.9919 | -35.5781 | -36.0898
25.998 | -35.5785 | -36.095
26.0042 | -35.5789 | -36.1004
26.0103 | -35.5792 | -36.1056
26.0164 | -35.5796 | -36.111
26.0225 | -35.58 | -36.1163
26.0286 | -35.5804 | -36.1219
26.0347 | -35.5808 | -36.1271
26.0408 | -35.5813 | -36.1327
26.0469 | -35.5817 | -36.138
26.053 | -35.5821 | -36.1437
26.0591 | -35.5825 | -36.1489
26.0652 | -35.5829 | -36.1541
26.0713 | -35.5832 | -36.1599
26.0774 | -35.5836 | -36.1654
26.0835 | -35.584 | -36.1709
26.0896 | -35.5844 | -36.1764
26.0957 | -35.5848 | -36.1818
26.1018 | -35.5853 | -36.1869
26.1079 | -35.5857 | -36.1923
26.114 | -35.586 | -36.1982
26.1201 | -35.5864 | -36.2037
26.1262 | -35.5868 | -36.2094
26.1323 | -35.5872 | -36.2145
26.1384 | -35.5876 | -36.2196
26.1445 | -35.5881 | -36.2249
26.1507 | -35.5884 | -36.2303
26.1568 | -35.5888 | -36.2358
26.1629 | -35.5892 | -36.2415
26.169 | -35.5896 | -36.2471
26.1751 | -35.59 | -36.2523
26.1812 | -35.5904 | -36.2578
26.1873 | -35.5908 | -36.2634
26.1934 | -35.5912 | -36.2688
26.1995 | -35.5916 | -36.2741
26.2056 | -35.592 | -36.2794
26.2117 | -35.5924 | -36.2851
26.2178 | -35.5928 | -36.2903
26.2239 | -35.5932 | -36.2958
26.23 | -35.5935 | -36.3014
26.2361 | -35.594 | -36.3071
26.2422 | -35.5943 | -36.3128
26.2483 | -35.5948 | -36.3182
26.2544 | -35.5951 | -36.3235
26.2605 | -35.5956 | -36.3291
26.2666 | -35.5959 | -36.3345
26.2727 | -35.5963 | -36.3399
26.2788 | -35.5967 | -36.345
26.2849 | -35.5971 | -36.3504
26.2911 | -35.5974 | -36.356
26.2972 | -35.5978 | -36.3613
26.3033 | -35.5982 | -36.3666
26.3094 | -35.5986 | -36.3723
26.3155 | -35.5989 | -36.3772
26.3216 | -35.5993 | -36.3829
26.3277 | -35.5997 | -36.3883
26.3338 | -35.6001 | -36.3937
26.3399 | -35.6005 | -36.3991
26.346 | -35.6009 | -36.4044
26.3521 | -35.6013 | -36.4101
26.3582 | -35.6017 | -36.4156
26.3643 | -35.6021 | -36.4211
26.3704 | -35.6025 | -36.4267
26.3765 | -35.603 | -36.4322
26.3826 | -35.6033 | -36.4376
26.3887 | -35.6037 | -36.4432
26.3948 | -35.6041 | -36.4486
26.4009 | -35.6045 | -36.4536
26.407 | -35.6049 | -36.4587
26.4131 | -35.6053 | -36.464
26.4192 | -35.6056 | -36.4691
26.4253 | -35.6061 | -36.4742
26.4314 | -35.6065 | -36.4796
26.4376 | -35.6069 | -36.4849
26.4437 | -35.6073 | -36.4903
26.4498 | -35.6077 | -36.4958
26.4559 | -35.6081 | -36.5013
26.462 | -35.6085 | -36.5065
26.4681 | -35.6089 | -36.5119
26.4742 | -35.6093 | -36.5175
26.4803 | -35.6097 | -36.5228
26.4864 | -35.6101 | -36.5282
26.4925 | -35.6104 | -36.5334
26.4986 | -35.6108 | -36.5389
26.5047 | -35.6112 | -36.5444
26.5108 | -35.6117 | -36.5497
26.5169 | -35.612 | -36.555
26.523 | -35.6124 | -36.5606
26.5291 | -35.6128 | -36.5657
26.5352 | -35.6132 | -36.5713
26.5413 | -35.6136 | -36.5767
26.5474 | -35.614 | -36.5819
26.5535 | -35.6144 | -36.5874
26.5596 | -35.6147 | -36.5926
26.5657 | -35.6151 | -36.5979
26.5718 | -35.6154 | -36.6033
26.578 | -35.6158 | -36.6086
26.5841 | -35.6162 | -36.614
26.5902 | -35.6166 | -36.6194
26.5963 | -35.617 | -36.6245
26.6024 | -35.6173 | -36.6299
26.6085 | -35.6177 | -36.6356
26.6146 | -35.6181 | -36.6407
26.6207 | -35.6185 | -36.6461
26.6268 | -35.6188 | -36.6517
26.6329 | -35.6192 | -36.6569
26.639 | -35.6196 | -36.6622
26.6451 | -35.6199 | -36.6677
26.6512 | -35.6203 | -36.6732
26.6573 | -35.6206 | -36.6786
26.6634 | -35.6209 | -36.6837
26.6695 | -35.6213 | -36.6892
26.6756 | -35.6217 | -36.6946
26.6817 | -35.6221 | -36.7001
26.6878 | -35.6224 | -36.7055
26.6939 | -35.6228 | -36.7111
26.7 | -35.6231 | -36.7162
26.7061 | -35.6235 | -36.7217
26.7122 | -35.6238 | -36.7273
26.7183 | -35.6242 | -36.7327
26.7245 | -35.6246 | -36.7381
26.7306 | -35.625 | -36.7433
26.7367 | -35.6254 | -36.7482
26.7428 | -35.6258 | -36.7533
26.7489 | -35.6261 | -36.7588
26.755 | -35.6265 | -36.7642
26.7611 | -35.6269 | -36.7697
26.7672 | -35.6273 | -36.7752
26.7733 | -35.6277 | -36.7804
26.7794 | -35.6281 | -36.7855
26.7855 | -35.6285 | -36.7908
26.7916 | -35.6288 | -36.7959
26.7977 | -35.6291 | -36.8012
26.8038 | -35.6295 | -36.8061
26.8099 | -35.6298 | -36.8118
26.816 | -35.6302 | -36.8171
26.8221 | -35.6305 | -36.8225
26.8282 | -35.6309 | -36.8278
26.8343 | -35.6313 | -36.8333
26.8404 | -35.6316 | -36.8386
26.8465 | -35.632 | -36.8439
26.8526 | -35.6324 | -36.8493
26.8587 | -35.6327 | -36.8546
26.8649 | -35.6331 | -36.86
26.871 | -35.6335 | -36.8652
26.8771 | -35.6339 | -36.8703
26.8832 | -35.6343 | -36.8754
26.8893 | -35.6347 | -36.8809
26.8954 | -35.6351 | -36.8859
26.9015 | -35.6355 | -36.8912
26.9076 | -35.6359 | -36.8968
26.9137 | -35.6363 | -36.9022
26.9198 | -35.6367 | -36.9077
26.9259 | -35.637 | -36.9128
26.932 | -35.6374 | -36.918
26.9381 | -35.6377 | -36.9237
26.9442 | -35.6381 | -36.9285
26.9503 | -35.6384 | -36.934
26.9564 | -35.6388 | -36.9395
26.9625 | -35.6391 | -36.9449
26.9686 | -35.6395 | -36.9504
26.9747 | -35.6399 | -36.9556
26.9808 | -35.6403 | -36.9607
26.9869 | -35.6406 | -36.966
26.993 | -35.6409 | -36.9713
26.9991 | -35.6413 | -36.9767
27.0052 | -35.6416 | -36.9823
27.0114 | -35.642 | -36.9876
27.0175 | -35.6423 | -36.9928
27.0236 | -35.6426 | -36.9981
27.0297 | -35.643 | -37.0037
27.0358 | -35.6433 | -37.009
27.0419 | -35.6437 | -37.0139
27.048 | -35.6441 | -37.0192
27.0541 | -35.6444 | -37.0246
27.0602 | -35.6448 | -37.03
27.0663 | -35.6452 | -37.0353
27.0724 | -35.6455 | -37.0403
27.0785 | -35.6459 | -37.0455
27.0846 | -35.6462 | -37.0508
27.0907 | -35.6465 | -37.0562
27.0968 | -35.6469 | -37.0618
27.1029 | -35.6473 | -37.0666
27.109 | -35.6476 | -37.0719
27.1151 | -35.648 | -37.0771
27.1212 | -35.6482 | -37.0819
27.1273 | -35.6486 | -37.0873
27.1334 | -35.6489 | -37.0925
27.1395 | -35.6493 | -37.0973
27.1456 | -35.6496 | -37.1023
27.1518 | -35.65 | -37.1076
27.1579 | -35.6503 | -37.1124
27.164 | -35.6507 | -37.1177
27.1701 | -35.651 | -37.123
27.1762 | -35.6514 | -37.1282
27.1823 | -35.6518 | -37.1332
27.1884 | -35.6522 | -37.1391
27.1945 | -35.6525 | -37.1441
27.2006 | -35.6528 | -37.1495
27.2067 | -35.6532 | -37.1546
27.2128 | -35.6535 | -37.1601
27.2189 | -35.6539 | -37.1656
27.225 | -35.6543 | -37.171
27.2311 | -35.6546 | -37.1762
27.2372 | -35.6549 | -37.1811
27.2433 | -35.6553 | -37.1862
27.2494 | -35.6556 | -37.1916
27.2555 | -35.656 | -37.1968
27.2616 | -35.6563 | -37.2014
27.2677 | -35.6567 | -37.2066
27.2738 | -35.6571 | -37.2115
27.2799 | -35.6574 | -37.2167
27.286 | -35.6578 | -37.2219
27.2921 | -35.6581 | -37.2273
27.2983 | -35.6585 | -37.2322
27.3044 | -35.6588 | -37.2373
27.3105 | -35.6592 | -37.2426
27.3166 | -35.6595 | -37.2478
27.3227 | -35.6599 | -37.2528
27.3288 | -35.6602 | -37.2582
27.3349 | -35.6606 | -37.2631
27.341 | -35.661 | -37.2686
27.3471 | -35.6613 | -37.2739
27.3532 | -35.6617 | -37.2787
27.3593 | -35.6621 | -37.2838
27.3654 | -35.6625 | -37.2891
27.3715 | -35.6628 | -37.2944
27.3776 | -35.6632 | -37.2995
27.3837 | -35.6636 | -37.3045
27.3898 | -35.6639 | -37.31
27.3959 | -35.6643 | -37.3151
27.402 | -35.6647 | -37.3199
27.4081 | -35.665 | -37.3252
27.4142 | -35.6654 | -37.3304
27.4203 | -35.6657 | -37.3355
27.4264 | -35.6661 | -37.3409
27.4325 | -35.6664 | -37.3457
27.4387 | -35.6667 | -37.3503
27.4448 | -35.667 | -37.3552
27.4509 | -35.6674 | -37.3606
27.457 | -35.6677 | -37.3656
27.4631 | -35.6681 | -37.3707
27.4692 | -35.6684 | -37.3761
27.4753 | -35.6686 | -37.3806
27.4814 | -35.6689 | -37.3856
27.4875 | -35.6693 | -37.3907
27.4936 | -35.6697 | -37.396
27.4997 | -35.6701 | -37.4012
27.5058 | -35.6704 | -37.4065
27.5119 | -35.6707 | -37.4114
27.518 | -35.6711 | -37.4165
27.5241 | -35.6715 | -37.4218
27.5302 | -35.6718 | -37.4272
27.5363 | -35.6721 | -37.4326
27.5424 | -35.6725 | -37.4377
27.5485 | -35.6727 | -37.4431
27.5546 | -35.6731 | -37.4481
27.5607 | -35.6734 | -37.4534
27.5668 | -35.6737 | -37.4585
27.5729 | -35.6741 | -37.4636
27.5791 | -35.6743 | -37.4688
27.5852 | -35.6747 | -37.474
27.5913 | -35.675 | -37.4795
27.5974 | -35.6753 | -37.4846
27.6035 | -35.6757 | -37.4899
27.6096 | -35.676 | -37.495
27.6157 | -35.6763 | -37.4999
27.6218 | -35.6766 | -37.505
27.6279 | -35.677 | -37.51
27.634 | -35.6773 | -37.5152
27.6401 | -35.6777 | -37.5205
27.6462 | -35.678 | -37.5257
27.6523 | -35.6784 | -37.5309
27.6584 | -35.6787 | -37.5362
27.6645 | -35.6791 | -37.5414
27.6706 | -35.6794 | -37.5468
27.6767 | -35.6798 | -37.5521
27.6828 | -35.6802 | -37.5571
27.6889 | -35.6805 | -37.5622
27.695 | -35.6808 | -37.5673
27.7011 | -35.6811 | -37.5722
27.7072 | -35.6815 | -37.5778
27.7133 | -35.6819 | -37.583
27.7194 | -35.6822 | -37.5878
27.7256 | -35.6825 | -37.5928
27.7317 | -35.6828 | -37.598
27.7378 | -35.6832 | -37.6031
27.7439 | -35.6835 | -37.608
27.75 | -35.6839 | -37.6129
27.7561 | -35.6842 | -37.6175
27.7622 | -35.6845 | -37.6223
27.7683 | -35.6848 | -37.6278
27.7744 | -35.6852 | -37.6329
27.7805 | -35.6855 | -37.638
27.7866 | -35.6859 | -37.6432
27.7927 | -35.6862 | -37.6485
27.7988 | -35.6865 | -37.6537
27.8049 | -35.6869 | -37.6585
27.811 | -35.6872 | -37.664
27.8171 | -35.6875 | -37.6693
27.8232 | -35.6877 | -37.6741
27.8293 | -35.688 | -37.6791
27.8354 | -35.6884 | -37.6843
27.8415 | -35.6887 | -37.6891
27.8476 | -35.689 | -37.6941
27.8537 | -35.6893 | -37.6991
27.8598 | -35.6896 | -37.7042
27.866 | -35.69 | -37.7093
27.8721 | -35.6903 | -37.7143
27.8782 | -35.6906 | -37.7194
27.8843 | -35.6909 | -37.7248
27.8904 | -35.6913 | -37.7304
27.8965 | -35.6916 | -37.7357
27.9026 | -35.6918 | -37.7405
27.9087 | -35.6922 | -37.7457
27.9148 | -35.6925 | -37.7506
27.9209 | -35.6928 | -37.756
27.927 | -35.6932 | -37.761
27.9331 | -35.6936 | -37.7659
27.9392 | -35.6939 | -37.7713
27.9453 | -35.6942 | -37.7763
27.9514 | -35.6946 | -37.7815
27.9575 | -35.6949 | -37.7868
27.9636 | -35.6952 | -37.7919
27.9697 | -35.6954 | -37.7971
27.9758 | -35.6958 | -37.8027
27.9819 | -35.6961 | -37.8073
27.988 | -35.6964 | -37.8122
27.9941 | -35.6967 | -37.8171
28.0002 | -35.6971 | -37.822
28.0063 | -35.6974 | -37.8273
28.0125 | -35.6977 | -37.8326
28.0186 | -35.698 | -37.8379
28.0247 | -35.6983 | -37.8431
28.0308 | -35.6987 | -37.8478
28.0369 | -35.699 | -37.8524
28.043 | -35.6994 | -37.8576
28.0491 | -35.6997 | -37.8629
28.0552 | -35.6999 | -37.8682
28.0613 | -35.7002 | -37.8735
28.0674 | -35.7005 | -37.8781
28.0735 | -35.7008 | -37.8833

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.5139 | -41.8171
21.4908 | -35.5422 | -34.9158
16.4948 | -34.9239 | -27.5587
12.6602 | -34.5646 | -21.2979
9.7171 | -34.3451 | -16.4089
7.4582 | -34.1963 | -12.4379
5.7244 | -34.0939 | -9.2124
4.3936 | -34.0251 | -6.6227
3.3722 | -34.0125 | -4.5766
2.5883 | -34.003 | -3.0246
1.9866 | -34.0022 | -1.7857
1.5248 | -34.0142 | -0.79191
1.1703 | -34.0104 | -0.1426
1 | -34.0268 | 0.17866
0.89822 | -34.0092 | 0.44994
0.68941 | -33.9902 | 1.0329
0.52919 | -33.9888 | 1.4842
0.40612 | -33.998 | 2.0909
0.31171 | -33.9933 | 2.5431
0.23928 | -34.0095 | 3.2659
0.18358 | -33.9985 | 3.6916
0.14091 | -34.0222 | 4.5702
0.10819 | -34.0554 | 5.7273
0.083017 | -34.0596 | 7.2157
0.063705 | -34.1227 | 9.0835
0.048923 | -34.1792 | 11.5777
0.037479 | -34.2879 | 14.7708
0.028801 | -34.4721 | 18.7972
0.022078 | -34.7779 | 24.2084
0.017023 | -35.2667 | 30.6772
0.013018 | -36.0736 | 38.3603
0.010014 | -37.1957 | 46.2704

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003628 | -1.2569
0.007256 | -24.5225
0.010884 | -32.1606
0.014512 | -37.6264
0.018141 | -39.5986
0.021769 | -42.9577
0.025397 | -45.2836
0.029025 | -45.9925
0.032653 | -48.0932
0.036281 | -49.5424
0.039909 | -50.8838
0.043537 | -52.4587
0.047166 | -53.6704
0.050794 | -54.6516
0.054422 | -55.4419
0.05805 | -56.9472
0.061678 | -58.0736
0.065306 | -57.7957
0.068934 | -57.5736
0.072562 | -56.7398
0.07619 | -58.256
0.079819 | -59.8715
0.083447 | -60.6333
0.087075 | -61.7162
0.090703 | -61.6662
0.094331 | -61.4931
0.097959 | -61.3491
0.10159 | -62.3075
0.10522 | -63.6847
0.10884 | -63.3405
0.11247 | -63.7801
0.1161 | -63.9595
0.11973 | -64.308
0.12336 | -64.8766
0.12698 | -64.7969
0.13061 | -64.5862
0.13424 | -65.718
0.13787 | -66.1325
0.1415 | -67.4575
0.14513 | -67.4932
0.14875 | -67.8568
0.15238 | -67.9595
0.15601 | -68.1616
0.15964 | -69.3675
0.16326 | -69.7215
0.16689 | -68.8194
0.17052 | -68.8028
0.17415 | -68.6414
0.17778 | -68.983
0.18141 | -68.7048
0.18503 | -68.9162
0.18866 | -68.4451
0.19229 | -69.1211
0.19592 | -69.7832
0.19955 | -69.5447
0.20317 | -69.2459
0.2068 | -69.9129
0.21043 | -71.0975
0.21406 | -71.5637
0.21769 | -72.1016
0.22132 | -71.1845
0.22494 | -71.0187
0.22857 | -72.054
0.2322 | -72.136
0.23583 | -72.7937
0.23946 | -73.3759
0.24308 | -73.7305
0.24671 | -73.9835
0.25034 | -73.317
0.25397 | -73.4275
0.2576 | -73.728
0.26122 | -73.1663
0.26485 | -73.3641
0.26848 | -73.7558
0.27211 | -74.1506
0.27574 | -75.1012
0.27936 | -75.7656
0.28299 | -75.2363
0.28662 | -75.5953
0.29025 | -74.9233
0.29388 | -74.159
0.29751 | -74.1712
0.30113 | -75.2259
0.30476 | -75.4659
0.30839 | -74.855
0.31202 | -74.7898
0.31565 | -74.9078
0.31927 | -75.0253
0.3229 | -75.3268
0.32653 | -76.0987
0.33016 | -76.5076
0.33379 | -75.4672
0.33742 | -75.9073
0.34104 | -75.3215
0.34467 | -76.0862
0.3483 | -77.1941
0.35193 | -77.2261
0.35556 | -78.2755
0.35918 | -78.7069
0.36281 | -78.5104
0.36644 | -78.63
0.37007 | -78.2389
0.3737 | -77.5614
0.37732 | -77.4027
0.38095 | -76.8483
0.38458 | -76.6803
0.38821 | -76.7852
0.39184 | -77.406
0.39547 | -77.5138
0.39909 | -77.149
0.40272 | -77.5312
0.40635 | -78.6747
0.40998 | -78.5421
0.41361 | -77.4318
0.41723 | -78.2864
0.42086 | -77.5941
0.42449 | -77.6015
0.42812 | -77.3945
0.43175 | -77.5468
0.43537 | -77.2821
0.439 | -78.9121
0.44263 | -78.6174
0.44626 | -77.957
0.44989 | -77.792
0.45352 | -78.1218
0.45714 | -78.3692
0.46077 | -78.1784
0.4644 | -77.9269
0.46803 | -78.8973
0.47165 | -78.534
0.47528 | -79.5824
0.47891 | -80.0625
0.48254 | -78.8383
0.48617 | -78.9329
0.4898 | -79.6251
0.49342 | -79.8685
0.49705 | -79.233
0.50068 | -79.2916
0.50431 | -79.6758
0.50794 | -80.5053
0.51157 | -80.8448
0.51519 | -80.2347
0.51882 | -80.2134
0.52245 | -79.9964
0.52608 | -80.3122
0.5297 | -80.7374
0.53333 | -80.5406
0.53696 | -79.7725
0.54059 | -79.2182
0.54422 | -79.6256
0.54785 | -80.528
0.55147 | -81.1013
0.5551 | -81.5559
0.55873 | -81.7779
0.56236 | -81.3501
0.56599 | -81.0549
0.56961 | -80.986
0.57324 | -80.7676
0.57687 | -80.8854
0.5805 | -81.15
0.58413 | -81.3672
0.58776 | -81.7771
0.59138 | -82.3049
0.59501 | -82.045
0.59864 | -81.5619
0.60227 | -81.7625
0.6059 | -82.4902
0.60952 | -82.7068
0.61315 | -82.3789
0.61678 | -82.0502
0.62041 | -81.9244
0.62404 | -81.9913
0.62766 | -82.1372
0.63129 | -82.3496
0.63492 | -82.4287
0.63855 | -82.3096
0.64218 | -82.1594
0.6458 | -82.0396
0.64943 | -82.2556
0.65306 | -82.3129
0.65669 | -82.0787
0.66032 | -81.9443
0.66395 | -81.7334
0.66757 | -81.8481
0.6712 | -82.2961
0.67483 | -82.7066
0.67846 | -82.8965
0.68209 | -82.8803
0.68571 | -82.9789
0.68934 | -82.8998
0.69297 | -82.7532
0.6966 | -82.9217
0.70023 | -83.5765
0.70386 | -83.7661
0.70748 | -83.4898
0.71111 | -83.6784
0.71474 | -83.5672
0.71837 | -83.2269
0.722 | -83.2168
0.72562 | -83.346
0.72925 | -83.2201
0.73288 | -82.8705
0.73651 | -82.6265
0.74014 | -82.935
0.74376 | -83.495
0.74739 | -83.7433
0.75102 | -83.9683
0.75465 | -84.1951
0.75828 | -84.2791
0.76191 | -84.0421
0.76553 | -84.0513
0.76916 | -84.1857
0.77279 | -83.7748
0.77642 | -82.9392
0.78004 | -82.245
0.78367 | -81.9339
0.7873 | -82.3453
0.79093 | -83.2704
0.79456 | -84.0302
0.79819 | -84.5124
0.80181 | -84.4187
0.80544 | -83.9572
0.80907 | -83.8968
0.8127 | -84.1532
0.81633 | -84.0942
0.81995 | -84.0549
0.82358 | -84.1773
0.82721 | -83.9816
0.83084 | -83.7681
0.83447 | -83.991
0.8381 | -84.3181
0.84172 | -84.407
0.84535 | -84.4847
0.84898 | -84.567
0.85261 | -84.6738
0.85624 | -84.8174
0.85986 | -85.4924
0.86349 | -85.9636
0.86712 | -85.9608
0.87075 | -85.9157
0.87438 | -85.5125
0.87801 | -85.1249
0.88163 | -85.1937
0.88526 | -85.4903
0.88889 | -85.4291
0.89252 | -85.1642
0.89614 | -85.4087
0.89977 | -85.7911
0.9034 | -85.4696
0.90703 | -85.304
0.91066 | -85.5477
0.91429 | -85.5482
0.91791 | -85.6245
0.92154 | -85.7905
0.92517 | -86.036
0.9288 | -85.9538
0.93243 | -85.6987
0.93605 | -85.8033
0.93968 | -86.1867
0.94331 | -86.2397
0.94694 | -85.5763
0.95057 | -85.3112
0.9542 | -85.6824
0.95782 | -86.6578
0.96145 | -87.7369
0.96508 | -87.7783
0.96871 | -87.008
0.97234 | -86.3598
0.97596 | -86.0847
0.97959 | -85.8997
0.98322 | -86.0371
0.98685 | -86.4547
0.99048 | -86.4065
0.9941 | -85.912
0.99773 | -85.724
1.0014 | -85.9726
1.005 | -86.329
1.0086 | -86.1968
1.0122 | -86.0631
1.0159 | -86.2408
1.0195 | -86.5178
1.0231 | -86.7873
1.0268 | -86.928
1.0304 | -86.6928
1.034 | -86.7414
1.0376 | -87.2061
1.0413 | -87.2794
1.0449 | -86.7953
1.0485 | -86.578
1.0522 | -86.6129
1.0558 | -86.8613
1.0594 | -87.097
1.063 | -87.1537
1.0667 | -87.1521
1.0703 | -87.1668
1.0739 | -87.0417
1.0776 | -86.7733
1.0812 | -86.543
1.0848 | -86.4523
1.0884 | -86.6119
1.0921 | -87.0135
1.0957 | -87.3076
1.0993 | -87.3257
1.1029 | -87.3758
1.1066 | -87.6115
1.1102 | -87.9182
1.1138 | -88.1631
1.1175 | -88.2184
1.1211 | -88.0272
1.1247 | -87.64
1.1283 | -87.3253
1.132 | -87.4271
1.1356 | -87.9914
1.1392 | -88.6292
1.1429 | -88.7723
1.1465 | -88.3435
1.1501 | -87.7414
1.1537 | -87.3248
1.1574 | -87.2228
1.161 | -87.3617
1.1646 | -87.6148
1.1683 | -87.8035
1.1719 | -87.847
1.1755 | -88.0035
1.1791 | -88.3647
1.1828 | -88.5898
1.1864 | -88.5015
1.19 | -88.2947
1.1937 | -88.1829
1.1973 | -88.287
1.2009 | -88.573
1.2045 | -88.819
1.2082 | -88.823
1.2118 | -88.6711
1.2154 | -88.5592
1.219 | -88.5192
1.2227 | -88.4751
1.2263 | -88.3669
1.2299 | -88.2401
1.2336 | -88.2188
1.2372 | -88.3844
1.2408 | -88.693
1.2444 | -88.9781
1.2481 | -89.0656
1.2517 | -88.9216
1.2553 | -88.6795
1.259 | -88.547
1.2626 | -88.53
1.2662 | -88.5063
1.2698 | -88.5428
1.2735 | -88.719
1.2771 | -88.8601
1.2807 | -88.7568
1.2844 | -88.424
1.288 | -88.0565
1.2916 | -87.9684
1.2952 | -88.3931
1.2989 | -89.1744
1.3025 | -89.7586
1.3061 | -89.806
1.3098 | -89.6189
1.3134 | -89.595
1.317 | -89.777
1.3206 | -89.9473
1.3243 | -89.871
1.3279 | -89.5944
1.3315 | -89.4322
1.3351 | -89.5733
1.3388 | -89.844
1.3424 | -89.9435
1.346 | -89.8066
1.3497 | -89.6079
1.3533 | -89.4352
1.3569 | -89.2816
1.3605 | -89.2851
1.3642 | -89.5858
1.3678 | -90.0979
1.3714 | -90.5627
1.3751 | -90.7045
1.3787 | -90.5309
1.3823 | -90.1601
1.3859 | -89.6654
1.3896 | -89.2683
1.3932 | -89.1244
1.3968 | -89.1668
1.4005 | -89.2469
1.4041 | -89.2679
1.4077 | -89.2595
1.4113 | -89.2869
1.415 | -89.4162
1.4186 | -89.5467
1.4222 | -89.6156
1.4259 | -89.7782
1.4295 | -89.9685
1.4331 | -89.8819
1.4367 | -89.5314
1.4404 | -89.229
1.444 | -89.2187
1.4476 | -89.4225
1.4512 | -89.6602
1.4549 | -89.8676
1.4585 | -89.9743
1.4621 | -89.9431
1.4658 | -89.9259
1.4694 | -90.0931
1.473 | -90.3895
1.4766 | -90.618
1.4803 | -90.7065
1.4839 | -90.6465
1.4875 | -90.4592
1.4912 | -90.3021
1.4948 | -90.3596
1.4984 | -90.6844
1.502 | -91.0492
1.5057 | -91.1268
1.5093 | -90.9288
1.5129 | -90.6416
1.5166 | -90.3761
1.5202 | -90.2795
1.5238 | -90.4135
1.5274 | -90.67
1.5311 | -90.9233
1.5347 | -91.0201
1.5383 | -90.869
1.5419 | -90.6343
1.5456 | -90.5306
1.5492 | -90.5869
1.5528 | -90.601
1.5565 | -90.4132
1.5601 | -90.2423
1.5637 | -90.3912
1.5673 | -90.6451
1.571 | -90.7785
1.5746 | -90.847
1.5782 | -90.8912
1.5819 | -90.9035
1.5855 | -90.8497
1.5891 | -90.8013
1.5927 | -90.8491
1.5964 | -90.9837
1.6 | -91.1085
1.6036 | -91.1787
1.6073 | -91.1823
1.6109 | -91.1107
1.6145 | -90.9986
1.6181 | -90.914
1.6218 | -90.8708
1.6254 | -90.8204
1.629 | -90.7664
1.6327 | -90.771
1.6363 | -90.8414
1.6399 | -90.9389
1.6435 | -91.0617
1.6472 | -91.233
1.6508 | -91.4126
1.6544 | -91.4737
1.6581 | -91.3888
1.6617 | -91.2809
1.6653 | -91.2524
1.6689 | -91.3106
1.6726 | -91.3689
1.6762 | -91.3647
1.6798 | -91.2911
1.6834 | -91.1702
1.6871 | -91.0509
1.6907 | -91.0152
1.6943 | -91.1306
1.698 | -91.37
1.7016 | -91.6207
1.7052 | -91.7901
1.7088 | -91.8542
1.7125 | -91.8737
1.7161 | -91.9339
1.7197 | -92.0888
1.7234 | -92.2963
1.727 | -92.4561
1.7306 | -92.4913
1.7342 | -92.3691
1.7379 | -92.1386
1.7415 | -91.9165
1.7451 | -91.8129
1.7488 | -91.8583
1.7524 | -92.0173
1.756 | -92.2433
1.7596 | -92.4693
1.7633 | -92.5822
1.7669 | -92.5526
1.7705 | -92.4648
1.7741 | -92.4055
1.7778 | -92.4053
1.7814 | -92.4449
1.785 | -92.4718
1.7887 | -92.4122
1.7923 | -92.2187
1.7959 | -91.9782
1.7995 | -91.8223
1.8032 | -91.8202
1.8068 | -91.9534
1.8104 | -92.1185
1.8141 | -92.2648
1.8177 | -92.3857
1.8213 | -92.4456
1.8249 | -92.37
1.8286 | -92.1458
1.8322 | -91.9335
1.8358 | -91.88
1.8395 | -91.9553
1.8431 | -92.065
1.8467 | -92.1304
1.8503 | -92.1061
1.854 | -92.0032
1.8576 | -91.917
1.8612 | -91.9553
1.8649 | -92.1276
1.8685 | -92.362
1.8721 | -92.5596
1.8757 | -92.6403
1.8794 | -92.6019
1.883 | -92.5308
1.8866 | -92.5248
1.8902 | -92.6309
1.8939 | -92.7882
1.8975 | -92.896
1.9011 | -92.9026
1.9048 | -92.8272
1.9084 | -92.7504
1.912 | -92.7358
1.9156 | -92.733
1.9193 | -92.6263
1.9229 | -92.4249
1.9265 | -92.2798
1.9302 | -92.3163
1.9338 | -92.5115
1.9374 | -92.7206
1.941 | -92.8168
1.9447 | -92.8077
1.9483 | -92.7771
1.9519 | -92.7871
1.9556 | -92.8021
1.9592 | -92.7528
1.9628 | -92.6749
1.9664 | -92.7074
1.9701 | -92.9132
1.9737 | -93.1539
1.9773 | -93.2358
1.981 | -93.1312
1.9846 | -93.0121
1.9882 | -93.0697
1.9918 | -93.2959
1.9955 | -93.4706
1.9991 | -93.3984
2.0027 | -93.127
2.0063 | -92.835
2.01 | -92.5891
2.0136 | -92.4073
2.0172 | -92.294
2.0209 | -92.2626
2.0245 | -92.333
2.0281 | -92.5198
2.0317 | -92.7684
2.0354 | -92.9502
2.039 | -92.9354
2.0426 | -92.7397
2.0463 | -92.51
2.0499 | -92.3865
2.0535 | -92.42
2.0571 | -92.6028
2.0608 | -92.886
2.0644 | -93.1486
2.068 | -93.2796
2.0717 | -93.3522
2.0753 | -93.4882
2.0789 | -93.6865
2.0825 | -93.817
2.0862 | -93.7999
2.0898 | -93.7054
2.0934 | -93.6695
2.0971 | -93.6427
2.1007 | -93.6276
2.1043 | -93.6126
2.1079 | -93.5905
2.1116 | -93.5747
2.1152 | -93.5647
2.1188 | -93.5703
2.1224 | -93.5619
2.1261 | -93.5212
2.1297 | -93.4646
2.1333 | -93.4099
2.137 | -93.3734
2.1406 | -93.3636
2.1442 | -93.3809
2.1478 | -93.4253
2.1515 | -93.498
2.1551 | -93.5682
2.1587 | -93.588
2.1624 | -93.556
2.166 | -93.4968
2.1696 | -93.4422
2.1732 | -93.4427
2.1769 | -93.5148
2.1805 | -93.6174
2.1841 | -93.681
2.1878 | -93.6512
2.1914 | -93.5464
2.195 | -93.4213
2.1986 | -93.3258
2.2023 | -93.2846
2.2059 | -93.3296
2.2095 | -93.4828
2.2132 | -93.6843
2.2168 | -93.8307
2.2204 | -93.8511
2.224 | -93.737
2.2277 | -93.5776
2.2313 | -93.4607
2.2349 | -93.437
2.2385 | -93.4966
2.2422 | -93.5809
2.2458 | -93.6518
2.2494 | -93.6813
2.2531 | -93.691
2.2567 | -93.7278
2.2603 | -93.8208
2.2639 | -93.9758
2.2676 | -94.1603
2.2712 | -94.319
2.2748 | -94.417
2.2785 | -94.4306
2.2821 | -94.372
2.2857 | -94.2643
2.2893 | -94.1593
2.293 | -94.0592
2.2966 | -93.9638
2.3002 | -93.8699
2.3039 | -93.779
2.3075 | -93.7126
2.3111 | -93.7093
2.3147 | -93.7718
2.3184 | -93.8683
2.322 | -93.9464
2.3256 | -93.953
2.3293 | -93.9001
2.3329 | -93.8605
2.3365 | -93.8762
2.3401 | -93.9648
2.3438 | -94.1062
2.3474 | -94.2768
2.351 | -94.4303
2.3546 | -94.5089
2.3583 | -94.4873
2.3619 | -94.4046
2.3655 | -94.3503
2.3692 | -94.367
2.3728 | -94.432
2.3764 | -94.4881
2.38 | -94.5077
2.3837 | -94.4523
2.3873 | -94.2894
2.3909 | -94.0904
2.3946 | -93.9457
2.3982 | -93.9112
2.4018 | -93.987
2.4054 | -94.1293
2.4091 | -94.2649
2.4127 | -94.3341
2.4163 | -94.3064
2.42 | -94.181
2.4236 | -94.0465
2.4272 | -93.9626
2.4308 | -93.9121
2.4345 | -93.8836
2.4381 | -93.8695
2.4417 | -93.8684
2.4454 | -93.8895
2.449 | -93.9286
2.4526 | -93.948
2.4562 | -93.9286
2.4599 | -93.8735
2.4635 | -93.8167
2.4671 | -93.826
2.4707 | -93.9424
2.4744 | -94.1513
2.478 | -94.4043
2.4816 | -94.6309
2.4853 | -94.7675
2.4889 | -94.8142
2.4925 | -94.7911
2.4961 | -94.6993
2.4998 | -94.5722
2.5034 | -94.4723
2.507 | -94.4616
2.5107 | -94.5483
2.5143 | -94.6825
2.5179 | -94.8249
2.5215 | -94.9682
2.5252 | -95.1075
2.5288 | -95.2186
2.5324 | -95.2822
2.5361 | -95.3069
2.5397 | -95.3003
2.5433 | -95.2526
2.5469 | -95.1707
2.5506 | -95.0923
2.5542 | -95.0481
2.5578 | -95.0169
2.5615 | -94.9798
2.5651 | -94.9269
2.5687 | -94.8518
2.5723 | -94.786
2.576 | -94.766
2.5796 | -94.7769
2.5832 | -94.7696
2.5868 | -94.6887
2.5905 | -94.5403
2.5941 | -94.3797
2.5977 | -94.2668
2.6014 | -94.2189
2.605 | -94.2401
2.6086 | -94.3258
2.6122 | -94.4179
2.6159 | -94.3726
2.6195 | -94.3208
2.6231 | -94.2308
2.6268 | -94.1335
2.6304 | -94.0948
2.634 | -94.1543
2.6376 | -94.32
2.6413 | -94.5766
2.6449 | -94.8867
2.6485 | -95.1886
2.6522 | -95.4096
2.6558 | -95.4877
2.6594 | -95.4068
2.663 | -95.2311
2.6667 | -95.0464
2.6703 | -94.9191
2.6739 | -94.8846
2.6776 | -94.9281
2.6812 | -95.0228
2.6848 | -95.1546
2.6884 | -95.2993
2.6921 | -95.4411
2.6957 | -95.591
2.6993 | -95.742
2.7029 | -95.8716
2.7066 | -95.9523
2.7102 | -95.9685
2.7138 | -95.9232
2.7175 | -95.8375
2.7211 | -95.7345
2.7247 | -95.6205
2.7283 | -95.5373
2.732 | -95.4952
2.7356 | -95.5153
2.7392 | -95.6041
2.7429 | -95.7412
2.7465 | -95.8883
2.7501 | -95.9788
2.7537 | -95.9722
2.7574 | -95.8598
2.761 | -95.6701
2.7646 | -95.4372
2.7683 | -95.2045
2.7719 | -95.0228
2.7755 | -94.9169
2.7791 | -94.9014
2.7828 | -94.9767
2.7864 | -95.123
2.79 | -95.3139
2.7937 | -95.5355
2.7973 | -95.7652
2.8009 | -95.9471
2.8045 | -96.0487
2.8082 | -96.0509
2.8118 | -95.9881
2.8154 | -95.9204
2.819 | -95.8735
2.8227 | -95.8652
2.8263 | -95.8836
2.8299 | -95.889
2.8336 | -95.8565
2.8372 | -95.784
2.8408 | -95.6679
2.8444 | -95.534
2.8481 | -95.4007
2.8517 | -95.2879
2.8553 | -95.2392
2.859 | -95.2578
2.8626 | -95.323
2.8662 | -95.4159
2.8698 | -95.4913
2.8735 | -95.5194
2.8771 | -95.4924
2.8807 | -95.4102
2.8844 | -95.2841
2.888 | -95.1554
2.8916 | -95.069
2.8952 | -95.0583
2.8989 | -95.1445
2.9025 | -95.2988
2.9061 | -95.5071
2.9098 | -95.7298
2.9134 | -95.9165
2.917 | -96.0391
2.9206 | -96.0918
2.9243 | -96.0732
2.9279 | -95.9922
2.9315 | -95.8803
2.9351 | -95.767
2.9388 | -95.6762
2.9424 | -95.6356
2.946 | -95.6546
2.9497 | -95.7313
2.9533 | -95.8437
2.9569 | -95.9733
2.9605 | -96.0722
2.9642 | -96.1195
2.9678 | -96.1113
2.9714 | -96.0513
2.9751 | -95.9468
2.9787 | -95.797
2.9823 | -95.6281
2.9859 | -95.4664
2.9896 | -95.3492
2.9932 | -95.2961
2.9968 | -95.319
3.0005 | -95.4107
3.0041 | -95.545
3.0077 | -95.7039
3.0113 | -95.8489
3.015 | -95.9707
3.0186 | -96.0686
3.0222 | -96.1389
3.0259 | -96.1853
3.0295 | -96.2084
3.0331 | -96.2205
3.0367 | -96.2279
3.0404 | -96.2223
3.044 | -96.1974
3.0476 | -96.148
3.0512 | -96.0923
3.0549 | -96.0431
3.0585 | -96.019
3.0621 | -96.0259
3.0658 | -96.0607
3.0694 | -96.1082
3.073 | -96.1301
3.0766 | -96.1055
3.0803 | -96.029
3.0839 | -95.906
3.0875 | -95.7812
3.0912 | -95.6807
3.0948 | -95.6227
3.0984 | -95.614
3.102 | -95.6661
3.1057 | -95.7568
3.1093 | -95.8593
3.1129 | -95.942
3.1166 | -95.9806
3.1202 | -95.9852
3.1238 | -95.9752
3.1274 | -95.9873
3.1311 | -96.0509
3.1347 | -96.1665
3.1383 | -96.3511
3.142 | -96.4953
3.1456 | -96.6254
3.1492 | -96.7207
3.1528 | -96.7612
3.1565 | -96.7435
3.1601 | -96.6763
3.1637 | -96.5688
3.1673 | -96.4497
3.171 | -96.3388
3.1746 | -96.2558
3.1782 | -96.2069
3.1819 | -96.1898
3.1855 | -96.1865
3.1891 | -96.1805
3.1927 | -96.1747
3.1964 | -96.1668
3.2 | -96.145
3.2036 | -96.1197
3.2073 | -96.1073
3.2109 | -96.1219
3.2145 | -96.1496
3.2181 | -96.1882
3.2218 | -96.2427
3.2254 | -96.3026
3.229 | -96.3584
3.2327 | -96.3921
3.2363 | -96.407
3.2399 | -96.4204
3.2435 | -96.443
3.2472 | -96.4745
3.2508 | -96.5162
3.2544 | -96.5657
3.258 | -96.6175
3.2617 | -96.6751
3.2653 | -96.7483
3.2689 | -96.8458
3.2726 | -96.9542
3.2762 | -97.0332
3.2798 | -97.1081
3.2834 | -97.1786
3.2871 | -97.2493
3.2907 | -97.3131
3.2943 | -97.3546
3.298 | -97.3603
3.3016 | -97.3422
3.3052 | -97.2938
3.3088 | -97.2288
3.3125 | -97.1498
3.3161 | -97.06
3.3197 | -96.9712
3.3234 | -96.8829
3.327 | -96.8203
3.3306 | -96.7962
3.3342 | -96.795
3.3379 | -96.819
3.3415 | -96.8606
3.3451 | -96.9034
3.3488 | -96.9318
3.3524 | -96.9412
3.356 | -96.9351
3.3596 | -96.9418
3.3633 | -96.9622
3.3669 | -97.01
3.3705 | -97.0792
3.3742 | -97.1477
3.3778 | -97.2052
3.3814 | -97.2348
3.385 | -97.2426
3.3887 | -97.2309
3.3923 | -97.2147
3.3959 | -97.1951
3.3995 | -97.1742
3.4032 | -97.1687
3.4068 | -97.1897
3.4104 | -97.2407
3.4141 | -97.3077
3.4177 | -97.395
3.4213 | -97.4867
3.4249 | -97.5606
3.4286 | -97.6046
3.4322 | -97.63
3.4358 | -97.6402
3.4395 | -97.6484
3.4431 | -97.6732
3.4467 | -97.6995
3.4503 | -97.7231
3.454 | -97.7347
3.4576 | -97.7202
3.4612 | -97.6769
3.4649 | -97.6005
3.4685 | -97.5113
3.4721 | -97.4009
3.4757 | -97.2853
3.4794 | -97.1961
3.483 | -97.1323
3.4866 | -97.1013
3.4902 | -97.1076
3.4939 | -97.1422
3.4975 | -97.1997
3.5011 | -97.2723
3.5048 | -97.3487
3.5084 | -97.4057
3.512 | -97.434
3.5156 | -97.439
3.5193 | -97.4376
3.5229 | -97.4381
3.5265 | -97.45
3.5302 | -97.4674
3.5338 | -97.4685
3.5374 | -97.4652
3.541 | -97.4528
3.5447 | -97.4395
3.5483 | -97.4287
3.5519 | -97.4036
3.5556 | -97.3691
3.5592 | -97.3345
3.5628 | -97.3209
3.5664 | -97.3379
3.5701 | -97.3812
3.5737 | -97.449
3.5773 | -97.5268
3.581 | -97.6003
3.5846 | -97.6604
3.5882 | -97.7103
3.5918 | -97.7322
3.5955 | -97.7157
3.5991 | -97.6656
3.6027 | -97.5801
3.6063 | -97.4704
3.61 | -97.3584
3.6136 | -97.2766
3.6172 | -97.2401
3.6209 | -97.2519
3.6245 | -97.313
3.6281 | -97.4105
3.6317 | -97.5178
3.6354 | -97.6161
3.639 | -97.6884
3.6426 | -97.7033
3.6463 | -97.6552
3.6499 | -97.5571
3.6535 | -97.4321
3.6571 | -97.3236
3.6608 | -97.3336
3.6644 | -97.3375
3.668 | -97.3864
3.6717 | -97.4564
3.6753 | -97.5537
3.6789 | -97.6605
3.6825 | -97.7636
3.6862 | -97.8553
3.6898 | -97.9504
3.6934 | -98.0418
3.6971 | -98.1342
3.7007 | -98.228
3.7043 | -98.3053
3.7079 | -98.3518
3.7116 | -98.3539
3.7152 | -98.3184
3.7188 | -98.2484
3.7224 | -98.1525
3.7261 | -98.0553
3.7297 | -97.9801
3.7333 | -97.95
3.737 | -97.9478
3.7406 | -97.9776
3.7442 | -98.028
3.7478 | -98.0827
3.7515 | -98.1133
3.7551 | -98.1137
3.7587 | -98.0762
3.7624 | -98.0159
3.766 | -97.9471
3.7696 | -97.8834
3.7732 | -97.8398
3.7769 | -97.8169
3.7805 | -97.8237
3.7841 | -97.8555
3.7878 | -97.895
3.7914 | -97.9388
3.795 | -97.9796
3.7986 | -98.0217
3.8023 | -98.0596
3.8059 | -98.084
3.8095 | -98.1003
3.8132 | -98.1103
3.8168 | -98.1041
3.8204 | -98.0752
3.824 | -98.028
3.8277 | -97.9715
3.8313 | -97.9103
3.8349 | -97.8511
3.8385 | -97.8021
3.8422 | -97.7688
3.8458 | -97.7534
3.8494 | -97.7594
3.8531 | -97.7738
3.8567 | -97.804
3.8603 | -97.8561
3.8639 | -97.9217
3.8676 | -97.9975
3.8712 | -98.0571
3.8748 | -98.0836
3.8785 | -98.0796
3.8821 | -98.0456
3.8857 | -97.987
3.8893 | -97.9037
3.893 | -97.812
3.8966 | -97.7277
3.9002 | -97.6544
3.9039 | -97.6086
3.9075 | -97.6088
3.9111 | -97.6701
3.9147 | -97.7922
3.9184 | -97.963
3.922 | -98.1608
3.9256 | -98.3598
3.9293 | -98.5186
3.9329 | -98.6191
3.9365 | -98.635
3.9401 | -98.5749
3.9438 | -98.4732
3.9474 | -98.3477
3.951 | -98.236
3.9546 | -98.1559
3.9583 | -98.108
3.9619 | -98.1128
3.9655 | -98.1571
3.9692 | -98.2242
3.9728 | -98.297
3.9764 | -98.3602
3.98 | -98.4
3.9837 | -98.4126
3.9873 | -98.4045
3.9909 | -98.3887
3.9946 | -98.37
3.9982 | -98.3663
4.0018 | -98.3732
4.0054 | -98.4056
4.0091 | -98.4609
4.0127 | -98.5281
4.0163 | -98.5832
4.02 | -98.6096
4.0236 | -98.5926
4.0272 | -98.5393
4.0308 | -98.4683
4.0345 | -98.3964
4.0381 | -98.3445
4.0417 | -98.3253
4.0454 | -98.3415
4.049 | -98.4049
4.0526 | -98.501
4.0562 | -98.6133
4.0599 | -98.7146
4.0635 | -98.7835
4.0671 | -98.798
4.0707 | -98.7834
4.0744 | -98.7379
4.078 | -98.6918
4.0816 | -98.6599
4.0853 | -98.6377
4.0889 | -98.6265
4.0925 | -98.6123
4.0961 | -98.5881
4.0998 | -98.5544
4.1034 | -98.505
4.107 | -98.4515
4.1107 | -98.4017
4.1143 | -98.3645
4.1179 | -98.3425
4.1215 | -98.3404
4.1252 | -98.3467
4.1288 | -98.3602
4.1324 | -98.3636
4.1361 | -98.3686
4.1397 | -98.3631
4.1433 | -98.3516
4.1469 | -98.3484
4.1506 | -98.3645
4.1542 | -98.4003
4.1578 | -98.4413
4.1615 | -98.4845
4.1651 | -98.5285
4.1687 | -98.5579
4.1723 | -98.5622
4.176 | -98.5552
4.1796 | -98.5362
4.1832 | -98.5025
4.1868 | -98.4924
4.1905 | -98.4847
4.1941 | -98.4846
4.1977 | -98.4963
4.2014 | -98.5159
4.205 | -98.5496
4.2086 | -98.586
4.2122 | -98.6128
4.2159 | -98.6287
4.2195 | -98.6436
4.2231 | -98.6578
4.2268 | -98.6631
4.2304 | -98.661
4.234 | -98.664
4.2376 | -98.6704
4.2413 | -98.6784
4.2449 | -98.6947
4.2485 | -98.7181
4.2522 | -98.7381
4.2558 | -98.7659
4.2594 | -98.7946
4.263 | -98.8252
4.2667 | -98.8539
4.2703 | -98.8841
4.2739 | -98.9105
4.2776 | -98.9398
4.2812 | -98.9716
4.2848 | -98.9921
4.2884 | -99.0059
4.2921 | -99.0231
4.2957 | -99.0385
4.2993 | -99.0665
4.3029 | -99.0989
4.3066 | -99.1352
4.3102 | -99.1666
4.3138 | -99.185
4.3175 | -99.1976
4.3211 | -99.1936
4.3247 | -99.16
4.3283 | -99.1063
4.332 | -99.0417
4.3356 | -98.9773
4.3392 | -98.9057
4.3429 | -98.8382
4.3465 | -98.7835
4.3501 | -98.7448
4.3537 | -98.74
4.3574 | -98.7649
4.361 | -98.8099
4.3646 | -98.878
4.3683 | -98.9507
4.3719 | -99.0222
4.3755 | -99.0867
4.3791 | -99.1411
4.3828 | -99.1885
4.3864 | -99.2364
4.39 | -99.2893
4.3937 | -99.3478
4.3973 | -99.3983
4.4009 | -99.4262
4.4045 | -99.4312
4.4082 | -99.4265
4.4118 | -99.4112
4.4154 | -99.3793
4.419 | -99.3415
4.4227 | -99.2894
4.4263 | -99.2243
4.4299 | -99.1534
4.4336 | -99.071
4.4372 | -98.9914
4.4408 | -98.9136
4.4444 | -98.8428
4.4481 | -98.7911
4.4517 | -98.7554
4.4553 | -98.7352
4.459 | -98.7349
4.4626 | -98.7545
4.4662 | -98.7942
4.4698 | -98.8427
4.4735 | -98.8962
4.4771 | -98.9565
4.4807 | -99.0179
4.4844 | -99.0841
4.488 | -99.1513
4.4916 | -99.2264
4.4952 | -99.2986
4.4989 | -99.3688
4.5025 | -99.4399
4.5061 | -99.5004
4.5098 | -99.5505
4.5134 | -99.5773
4.517 | -99.5805
4.5206 | -99.5703
4.5243 | -99.5456
4.5279 | -99.5127
4.5315 | -99.4696
4.5351 | -99.4205
4.5388 | -99.3748
4.5424 | -99.3272
4.546 | -99.2719
4.5497 | -99.2218
4.5533 | -99.182
4.5569 | -99.1608
4.5605 | -99.1603
4.5642 | -99.1938
4.5678 | -99.247
4.5714 | -99.3149
4.5751 | -99.3987
4.5787 | -99.4887
4.5823 | -99.5741
4.5859 | -99.6423
4.5896 | -99.6851
4.5932 | -99.698
4.5968 | -99.6838
4.6005 | -99.6443
4.6041 | -99.5796
4.6077 | -99.4985
4.6113 | -99.4202
4.615 | -99.3525
4.6186 | -99.2998
4.6222 | -99.2603
4.6258 | -99.2309
4.6295 | -99.2159
4.6331 | -99.2084
4.6367 | -99.2075
4.6404 | -99.2092
4.644 | -99.2126
4.6476 | -99.2101
4.6512 | -99.2102
4.6549 | -99.1941
4.6585 | -99.1771
4.6621 | -99.1689
4.6658 | -99.1661
4.6694 | -99.1697
4.673 | -99.1842
4.6766 | -99.2147
4.6803 | -99.2518
4.6839 | -99.2977
4.6875 | -99.3401
4.6912 | -99.3861
4.6948 | -99.4437
4.6984 | -99.5095
4.702 | -99.5847
4.7057 | -99.6632
4.7093 | -99.7306
4.7129 | -99.7957
4.7166 | -99.8465
4.7202 | -99.8795
4.7238 | -99.8863
4.7274 | -99.8863
4.7311 | -99.8694
4.7347 | -99.834
4.7383 | -99.8048
4.742 | -99.7919
4.7456 | -99.7964
4.7492 | -99.8183
4.7528 | -99.8592
4.7565 | -99.9137
4.7601 | -99.9785
4.7637 | -100.0389
4.7673 | -100.0869
4.771 | -100.1213
4.7746 | -100.143
4.7782 | -100.1559
4.7819 | -100.1491
4.7855 | -100.1298
4.7891 | -100.0993
4.7927 | -100.0631
4.7964 | -100.0276
4.8 | -99.9871
4.8036 | -99.951
4.8073 | -99.9243
4.8109 | -99.9082
4.8145 | -99.896
4.8181 | -99.8867
4.8218 | -99.8777
4.8254 | -99.8703
4.829 | -99.8607
4.8327 | -99.8497
4.8363 | -99.838
4.8399 | -99.8156
4.8435 | -99.7851
4.8472 | -99.7556
4.8508 | -99.7206
4.8544 | -99.678
4.8581 | -99.6368
4.8617 | -99.5951
4.8653 | -99.5626
4.8689 | -99.5377
4.8726 | -99.5273
4.8762 | -99.5412
4.8798 | -99.5689
4.8834 | -99.6139
4.8871 | -99.6768
4.8907 | -99.745
4.8943 | -99.8139
4.898 | -99.8898
4.9016 | -99.959
4.9052 | -100.0172
4.9088 | -100.0634
4.9125 | -100.0907
4.9161 | -100.1032
4.9197 | -100.0928
4.9234 | -100.0551
4.927 | -100.0065
4.9306 | -99.9563
4.9342 | -99.9065
4.9379 | -99.8649
4.9415 | -99.8374
4.9451 | -99.833
4.9488 | -99.8443
4.9524 | -99.8615
4.956 | -99.8862
4.9596 | -99.915
4.9633 | -99.9376
4.9669 | -99.952
4.9705 | -99.9683
4.9741 | -99.9789
4.9778 | -99.9815
4.9814 | -99.9826
4.985 | -99.9757
4.9887 | -99.9656
4.9923 | -99.9467
4.9959 | -99.9271
4.9995 | -99.9154
5.0032 | -99.9097
5.0068 | -99.9018
5.0104 | -99.8971
5.0141 | -99.9001
5.0177 | -99.8992
5.0213 | -99.8952
5.0249 | -99.8908
5.0286 | -99.8866
5.0322 | -99.8881
5.0358 | -99.901
5.0395 | -99.9305
5.0431 | -99.9764
5.0467 | -100.0346
5.0503 | -100.1017
5.054 | -100.1746
5.0576 | -100.2435
5.0612 | -100.3002
5.0649 | -100.3358
5.0685 | -100.3554
5.0721 | -100.3563
5.0757 | -100.3387
5.0794 | -100.318
5.083 | -100.2964
5.0866 | -100.2652
5.0902 | -100.218
5.0939 | -100.1661
5.0975 | -100.1119
5.1011 | -100.0532
5.1048 | -100.007
5.1084 | -99.9663
5.112 | -99.9312
5.1156 | -99.9124
5.1193 | -99.9117
5.1229 | -99.9208
5.1265 | -99.9428
5.1302 | -99.9789
5.1338 | -100.0146
5.1374 | -100.0535
5.141 | -100.1028
5.1447 | -100.1509
5.1483 | -100.1915
5.1519 | -100.2316
5.1556 | -100.2714
5.1592 | -100.3207
5.1628 | -100.3779
5.1664 | -100.4388
5.1701 | -100.5013
5.1737 | -100.5636
5.1773 | -100.6241
5.181 | -100.6641
5.1846 | -100.6881
5.1882 | -100.686
5.1918 | -100.6519
5.1955 | -100.5936
5.1991 | -100.5184
5.2027 | -100.4435
5.2063 | -100.3751
5.21 | -100.321
5.2136 | -100.2838
5.2172 | -100.2675
5.2209 | -100.264
5.2245 | -100.2672
5.2281 | -100.2638
5.2317 | -100.2503
5.2354 | -100.2334
5.239 | -100.2122
5.2426 | -100.1957
5.2463 | -100.1843
5.2499 | -100.1737
5.2535 | -100.1614
5.2571 | -100.1502
5.2608 | -100.1476
5.2644 | -100.1507
5.268 | -100.1585
5.2717 | -100.1761
5.2753 | -100.2064
5.2789 | -100.2521
5.2825 | -100.3004
5.2862 | -100.3528
5.2898 | -100.4158
5.2934 | -100.4683
5.2971 | -100.5122
5.3007 | -100.5594
5.3043 | -100.5956
5.3079 | -100.6292
5.3116 | -100.6487
5.3152 | -100.6566
5.3188 | -100.6491
5.3224 | -100.6296
5.3261 | -100.5939
5.3297 | -100.5489
5.3333 | -100.4962
5.337 | -100.4451
5.3406 | -100.395
5.3442 | -100.3487
5.3478 | -100.3108
5.3515 | -100.2834
5.3551 | -100.2632
5.3587 | -100.2556
5.3624 | -100.2588
5.366 | -100.2774
5.3696 | -100.308
5.3732 | -100.3418
5.3769 | -100.3819
5.3805 | -100.419
5.3841 | -100.4461
5.3878 | -100.4696
5.3914 | -100.4765
5.395 | -100.4789
5.3986 | -100.4721
5.4023 | -100.4604
5.4059 | -100.4483
5.4095 | -100.4334
5.4132 | -100.4207
5.4168 | -100.4171
5.4204 | -100.4186
5.424 | -100.4229
5.4277 | -100.4277
5.4313 | -100.4368
5.4349 | -100.4478
5.4385 | -100.4577
5.4422 | -100.4589
5.4458 | -100.46
5.4494 | -100.4597
5.4531 | -100.4584
5.4567 | -100.4622
5.4603 | -100.47
5.4639 | -100.478
5.4676 | -100.486
5.4712 | -100.498
5.4748 | -100.5067
5.4785 | -100.5116
5.4821 | -100.5207
5.4857 | -100.5185
5.4893 | -100.4979
5.493 | -100.4683
5.4966 | -100.4246
5.5002 | -100.3591
5.5039 | -100.2845
5.5075 | -100.2113
5.5111 | -100.1366
5.5147 | -100.0697
5.5184 | -100.0144
5.522 | -99.9674
5.5256 | -99.9416
5.5293 | -99.9265
5.5329 | -99.9237
5.5365 | -99.9314
5.5401 | -99.9401
5.5438 | -99.962
5.5474 | -99.9907
5.551 | -100.0295
5.5546 | -100.0743
5.5583 | -100.1273
5.5619 | -100.1832
5.5655 | -100.2268
5.5692 | -100.2585
5.5728 | -100.2778
5.5764 | -100.2825
5.58 | -100.2698
5.5837 | -100.2401
5.5873 | -100.1825
5.5909 | -100.1148
5.5946 | -100.0493
5.5982 | -99.9785
5.6018 | -99.9112
5.6054 | -99.8516
5.6091 | -99.7919
5.6127 | -99.7406
5.6163 | -99.692
5.62 | -99.6412
5.6236 | -99.5956
5.6272 | -99.5446
5.6308 | -99.4923
5.6345 | -99.4358
5.6381 | -99.376
5.6417 | -99.3202
5.6454 | -99.27
5.649 | -99.2328
5.6526 | -99.2065
5.6562 | -99.1807
5.6599 | -99.1567
5.6635 | -99.1412
5.6671 | -99.1384
5.6707 | -99.1475
5.6744 | -99.1543
5.678 | -99.1635
5.6816 | -99.1781
5.6853 | -99.191
5.6889 | -99.2041
5.6925 | -99.2208
5.6961 | -99.2402
5.6998 | -99.2645
5.7034 | -99.2959
5.707 | -99.3247
5.7107 | -99.3563
5.7143 | -99.3878
5.7179 | -99.4115
5.7215 | -99.4231
5.7252 | -99.4198
5.7288 | -99.3987
5.7324 | -99.3583
5.7361 | -99.2996
5.7397 | -99.2385
5.7433 | -99.1727
5.7469 | -99.0991
5.7506 | -99.025
5.7542 | -98.9542
5.7578 | -98.8776
5.7615 | -98.8106
5.7651 | -98.7424
5.7687 | -98.6731
5.7723 | -98.6119
5.776 | -98.5508
5.7796 | -98.4909
5.7832 | -98.4302
5.7868 | -98.3781
5.7905 | -98.3356
5.7941 | -98.3049
5.7977 | -98.2805
5.8014 | -98.2709
5.805 | -98.2715
5.8086 | -98.2831
5.8122 | -98.3109
5.8159 | -98.3463
5.8195 | -98.3933
5.8231 | -98.4482
5.8268 | -98.5032
5.8304 | -98.5592
5.834 | -98.617
5.8376 | -98.6668
5.8413 | -98.7071
5.8449 | -98.7368
5.8485 | -98.7474
5.8522 | -98.7316
5.8558 | -98.7006
5.8594 | -98.6621
5.863 | -98.6089
5.8667 | -98.5545
5.8703 | -98.5
5.8739 | -98.4472
5.8776 | -98.3982
5.8812 | -98.3603
5.8848 | -98.333
5.8884 | -98.3162
5.8921 | -98.306
5.8957 | -98.3064
5.8993 | -98.3112
5.9029 | -98.3163
5.9066 | -98.3302
5.9102 | -98.3466
5.9138 | -98.3675
5.9175 | -98.3917
5.9211 | -98.4165
5.9247 | -98.4396
5.9283 | -98.4662
5.932 | -98.4937
5.9356 | -98.5201
5.9392 | -98.55
5.9429 | -98.5868
5.9465 | -98.6263
5.9501 | -98.6642
5.9537 | -98.7037
5.9574 | -98.7426
5.961 | -98.7844
5.9646 | -98.8316
5.9683 | -98.8755
5.9719 | -98.9126
5.9755 | -98.9557
5.9791 | -99.0084
5.9828 | -99.067
5.9864 | -99.1348
5.99 | -99.2069
5.9937 | -99.2783
5.9973 | -99.3521
6.0009 | -99.4233
6.0045 | -99.4861
6.0082 | -99.5393
6.0118 | -99.5907
6.0154 | -99.6432
6.019 | -99.6864
6.0227 | -99.7311
6.0263 | -99.7763
6.0299 | -99.8177
6.0336 | -99.8542
6.0372 | -99.8961
6.0408 | -99.945
6.0444 | -100.0016
6.0481 | -100.0639
6.0517 | -100.1269
6.0553 | -100.1971
6.059 | -100.2715
6.0626 | -100.3537
6.0662 | -100.4307
6.0698 | -100.5025
6.0735 | -100.5706
6.0771 | -100.6252
6.0807 | -100.6704
6.0844 | -100.7116
6.088 | -100.7417
6.0916 | -100.759
6.0952 | -100.7717
6.0989 | -100.7733
6.1025 | -100.7613
6.1061 | -100.7423
6.1098 | -100.7292
6.1134 | -100.718
6.117 | -100.7174
6.1206 | -100.7297
6.1243 | -100.7553
6.1279 | -100.79
6.1315 | -100.8348
6.1351 | -100.8846
6.1388 | -100.9406
6.1424 | -101.0025
6.146 | -101.0613
6.1497 | -101.1186
6.1533 | -101.1696
6.1569 | -101.207
6.1605 | -101.234
6.1642 | -101.2548
6.1678 | -101.2829
6.1714 | -101.3062
6.1751 | -101.3175
6.1787 | -101.3264
6.1823 | -101.3367
6.1859 | -101.3488
6.1896 | -101.3699
6.1932 | -101.3953
6.1968 | -101.4222
6.2005 | -101.456
6.2041 | -101.4997
6.2077 | -101.5486
6.2113 | -101.5991
6.215 | -101.6447
6.2186 | -101.6857
6.2222 | -101.7162
6.2259 | -101.7384
6.2295 | -101.7416
6.2331 | -101.7313
6.2367 | -101.712
6.2404 | -101.6822
6.244 | -101.6545
6.2476 | -101.6454
6.2512 | -101.6466
6.2549 | -101.6486
6.2585 | -101.6534
6.2621 | -101.6715
6.2658 | -101.7043
6.2694 | -101.7439
6.273 | -101.7904
6.2766 | -101.8371
6.2803 | -101.8802
6.2839 | -101.921
6.2875 | -101.9592
6.2912 | -101.9908
6.2948 | -102.0111
6.2984 | -102.0236
6.302 | -102.0296
6.3057 | -102.0252
6.3093 | -102.0095
6.3129 | -101.981
6.3166 | -101.946
6.3202 | -101.9021
6.3238 | -101.8545
6.3274 | -101.8062
6.3311 | -101.7598
6.3347 | -101.7121
6.3383 | -101.6685
6.3419 | -101.6315
6.3456 | -101.6013
6.3492 | -101.5815
6.3528 | -101.5666
6.3565 | -101.5547
6.3601 | -101.5449
6.3637 | -101.5339
6.3673 | -101.5243
6.371 | -101.5152
6.3746 | -101.5076
6.3782 | -101.4982
6.3819 | -101.4979
6.3855 | -101.5013
6.3891 | -101.5138
6.3927 | -101.5365
6.3964 | -101.562
6.4 | -101.5945
6.4036 | -101.6364
6.4073 | -101.6941
6.4109 | -101.7539
6.4145 | -101.8083
6.4181 | -101.8667
6.4218 | -101.9185
6.4254 | -101.9615
6.429 | -101.9976
6.4327 | -102.0246
6.4363 | -102.03
6.4399 | -102.0199
6.4435 | -101.9971
6.4472 | -101.9623
6.4508 | -101.9155
6.4544 | -101.8623
6.4581 | -101.8046
6.4617 | -101.7518
6.4653 | -101.7033
6.4689 | -101.6656
6.4726 | -101.641
6.4762 | -101.6297
6.4798 | -101.6307
6.4834 | -101.6483
6.4871 | -101.6739
6.4907 | -101.7045
6.4943 | -101.7486
6.498 | -101.7986
6.5016 | -101.8513
6.5052 | -101.9089
6.5088 | -101.9704
6.5125 | -102.0331
6.5161 | -102.0928
6.5197 | -102.1547
6.5234 | -102.2114
6.527 | -102.2585
6.5306 | -102.2983
6.5342 | -102.3329
6.5379 | -102.3574
6.5415 | -102.3835
6.5451 | -102.4062
6.5488 | -102.4251
6.5524 | -102.4429
6.556 | -102.4581
6.5596 | -102.4705
6.5633 | -102.4802
6.5669 | -102.483
6.5705 | -102.4824
6.5742 | -102.484
6.5778 | -102.4796
6.5814 | -102.4764
6.585 | -102.4805
6.5887 | -102.4843
6.5923 | -102.4826
6.5959 | -102.4845
6.5995 | -102.489
6.6032 | -102.4916
6.6068 | -102.5033
6.6104 | -102.5192
6.6141 | -102.5202
6.6177 | -102.523
6.6213 | -102.5238
6.6249 | -102.5175
6.6286 | -102.5062
6.6322 | -102.4868
6.6358 | -102.4586
6.6395 | -102.423
6.6431 | -102.3802
6.6467 | -102.3331
6.6503 | -102.2857
6.654 | -102.2402
6.6576 | -102.1982
6.6612 | -102.1572
6.6649 | -102.1224
6.6685 | -102.0952
6.6721 | -102.0701
6.6757 | -102.0548
6.6794 | -102.0457
6.683 | -102.0415
6.6866 | -102.0497
6.6902 | -102.0672
6.6939 | -102.0831
6.6975 | -102.1037
6.7011 | -102.1261
6.7048 | -102.1453
6.7084 | -102.1631
6.712 | -102.1861
6.7156 | -102.1985
6.7193 | -102.2068
6.7229 | -102.2169
6.7265 | -102.2311
6.7302 | -102.2432
6.7338 | -102.2516
6.7374 | -102.2533
6.741 | -102.2624
6.7447 | -102.2705
6.7483 | -102.2706
6.7519 | -102.2701
6.7556 | -102.2701
6.7592 | -102.2746
6.7628 | -102.2788
6.7664 | -102.2805
6.7701 | -102.2791
6.7737 | -102.2821
6.7773 | -102.2892
6.781 | -102.2976
6.7846 | -102.3116
6.7882 | -102.3306
6.7918 | -102.3594
6.7955 | -102.3881
6.7991 | -102.4207
6.8027 | -102.4496
6.8063 | -102.4767
6.81 | -102.5079
6.8136 | -102.5376
6.8172 | -102.5657
6.8209 | -102.5876
6.8245 | -102.607
6.8281 | -102.6245
6.8317 | -102.6397
6.8354 | -102.6544
6.839 | -102.6676
6.8426 | -102.6694
6.8463 | -102.6613
6.8499 | -102.6486
6.8535 | -102.6261
6.8571 | -102.6004
6.8608 | -102.5732
6.8644 | -102.5496
6.868 | -102.5281
6.8717 | -102.5116
6.8753 | -102.4949
6.8789 | -102.4757
6.8825 | -102.4538
6.8862 | -102.4363
6.8898 | -102.4168
6.8934 | -102.3936
6.8971 | -102.367
6.9007 | -102.3389
6.9043 | -102.3127
6.9079 | -102.2819
6.9116 | -102.2504
6.9152 | -102.2184
6.9188 | -102.1918
6.9224 | -102.1689
6.9261 | -102.1541
6.9297 | -102.1452
6.9333 | -102.1397
6.937 | -102.1463
6.9406 | -102.1621
6.9442 | -102.1868
6.9478 | -102.2125
6.9515 | -102.2485
6.9551 | -102.2925
6.9587 | -102.3395
6.9624 | -102.3824
6.966 | -102.416
6.9696 | -102.4518
6.9732 | -102.4902
6.9769 | -102.5208
6.9805 | -102.5383
6.9841 | -102.5476
6.9878 | -102.5565
6.9914 | -102.5462
6.995 | -102.5246
6.9986 | -102.5057
7.0023 | -102.4881
7.0059 | -102.468
7.0095 | -102.4485
7.0132 | -102.4302
7.0168 | -102.4133
7.0204 | -102.404
7.024 | -102.3943
7.0277 | -102.386
7.0313 | -102.3687
7.0349 | -102.3522
7.0385 | -102.3333
7.0422 | -102.3044
7.0458 | -102.2765
7.0494 | -102.2451
7.0531 | -102.2146
7.0567 | -102.178
7.0603 | -102.1409
7.0639 | -102.1113
7.0676 | -102.0911
7.0712 | -102.0834
7.0748 | -102.0746
7.0785 | -102.0735
7.0821 | -102.0894
7.0857 | -102.1184
7.0893 | -102.1616
7.093 | -102.2129
7.0966 | -102.2766
7.1002 | -102.3525
7.1039 | -102.4318
7.1075 | -102.5118
7.1111 | -102.5994
7.1147 | -102.688
7.1184 | -102.7675
7.122 | -102.8335
7.1256 | -102.8834
7.1293 | -102.926
7.1329 | -102.9555
7.1365 | -102.9748
7.1401 | -102.9632
7.1438 | -102.9373
7.1474 | -102.9086
7.151 | -102.8659
7.1546 | -102.8148
7.1583 | -102.7639
7.1619 | -102.7092
7.1655 | -102.6536
7.1692 | -102.5973
7.1728 | -102.5405
7.1764 | -102.491
7.18 | -102.4475
7.1837 | -102.4
7.1873 | -102.3548
7.1909 | -102.318
7.1946 | -102.2845
7.1982 | -102.2555
7.2018 | -102.2389
7.2054 | -102.2286
7.2091 | -102.2235
7.2127 | -102.2297
7.2163 | -102.2422
7.22 | -102.263
7.2236 | -102.2974
7.2272 | -102.3391
7.2308 | -102.3817
7.2345 | -102.425
7.2381 | -102.4678
7.2417 | -102.5047
7.2454 | -102.5421
7.249 | -102.5723
7.2526 | -102.5894
7.2562 | -102.5951
7.2599 | -102.5945
7.2635 | -102.5888
7.2671 | -102.5786
7.2707 | -102.5607
7.2744 | -102.5355
7.278 | -102.5181
7.2816 | -102.503
7.2853 | -102.4922
7.2889 | -102.4903
7.2925 | -102.4945
7.2961 | -102.5018
7.2998 | -102.5136
7.3034 | -102.5286
7.307 | -102.546
7.3107 | -102.5702
7.3143 | -102.5984
7.3179 | -102.6285
7.3215 | -102.6615
7.3252 | -102.707
7.3288 | -102.7506
7.3324 | -102.7897
7.3361 | -102.8332
7.3397 | -102.8749
7.3433 | -102.9178
7.3469 | -102.9611
7.3506 | -103.0019
7.3542 | -103.0369
7.3578 | -103.0597
7.3615 | -103.0815
7.3651 | -103.1026
7.3687 | -103.1126
7.3723 | -103.1204
7.376 | -103.1294
7.3796 | -103.1332
7.3832 | -103.1317
7.3868 | -103.117
7.3905 | -103.1025
7.3941 | -103.0889
7.3977 | -103.0728
7.4014 | -103.0421
7.405 | -103.0014
7.4086 | -102.9588
7.4122 | -102.9182
7.4159 | -102.8785
7.4195 | -102.8393
7.4231 | -102.8048
7.4268 | -102.7734
7.4304 | -102.746
7.434 | -102.7274
7.4376 | -102.7147
7.4413 | -102.7086
7.4449 | -102.7171
7.4485 | -102.7283
7.4522 | -102.7472
7.4558 | -102.7681
7.4594 | -102.7952
7.463 | -102.8227
7.4667 | -102.8495
7.4703 | -102.8769
7.4739 | -102.8993
7.4776 | -102.9193
7.4812 | -102.9439
7.4848 | -102.9603
7.4884 | -102.9741
7.4921 | -102.9919
7.4957 | -103.0112
7.4993 | -103.0275
7.5029 | -103.0362
7.5066 | -103.0472
7.5102 | -103.0613
7.5138 | -103.0727
7.5175 | -103.0813
7.5211 | -103.0845
7.5247 | -103.0802
7.5283 | -103.0708
7.532 | -103.0598
7.5356 | -103.0451
7.5392 | -103.0324
7.5429 | -103.0144
7.5465 | -102.9951
7.5501 | -102.9818
7.5537 | -102.9677
7.5574 | -102.9504
7.561 | -102.9389
7.5646 | -102.9323
7.5683 | -102.9258
7.5719 | -102.9139
7.5755 | -102.9146
7.5791 | -102.9204
7.5828 | -102.9321
7.5864 | -102.9476
7.59 | -102.9658
7.5937 | -102.9827
7.5973 | -103.0026
7.6009 | -103.0187
7.6045 | -103.0334
7.6082 | -103.0528
7.6118 | -103.0726
7.6154 | -103.0918
7.619 | -103.1068
7.6227 | -103.1187
7.6263 | -103.1377
7.6299 | -103.1505
7.6336 | -103.1572
7.6372 | -103.1618
7.6408 | -103.1649
7.6444 | -103.1656
7.6481 | -103.164
7.6517 | -103.158
7.6553 | -103.1557
7.659 | -103.1481
7.6626 | -103.1355
7.6662 | -103.1271
7.6698 | -103.1192
7.6735 | -103.1056
7.6771 | -103.0876
7.6807 | -103.0688
7.6844 | -103.0531
7.688 | -103.0301
7.6916 | -103.0088
7.6952 | -102.9898
7.6989 | -102.9762
7.7025 | -102.9704
7.7061 | -102.9723
7.7098 | -102.9854
7.7134 | -103.0121
7.717 | -103.0422
7.7206 | -103.0731
7.7243 | -103.1121
7.7279 | -103.1556
7.7315 | -103.1984
7.7351 | -103.2373
7.7388 | -103.2743
7.7424 | -103.31
7.746 | -103.344
7.7497 | -103.3675
7.7533 | -103.3866
7.7569 | -103.3993
7.7605 | -103.4035
7.7642 | -103.3974
7.7678 | -103.3847
7.7714 | -103.372
7.7751 | -103.3504
7.7787 | -103.3229
7.7823 | -103.2921
7.7859 | -103.2639
7.7896 | -103.235
7.7932 | -103.2072
7.7968 | -103.1828
7.8005 | -103.1588
7.8041 | -103.1341
7.8077 | -103.1119
7.8113 | -103.093
7.815 | -103.0825
7.8186 | -103.076
7.8222 | -103.0766
7.8258 | -103.0786
7.8295 | -103.0921
7.8331 | -103.1177
7.8367 | -103.1489
7.8404 | -103.1859
7.844 | -103.2308
7.8476 | -103.2769
7.8512 | -103.3032
7.8549 | -103.3325
7.8585 | -103.3582
7.8621 | -103.3733
7.8658 | -103.3883
7.8694 | -103.3889
7.873 | -103.388
7.8766 | -103.3897
7.8803 | -103.3877
7.8839 | -103.3777
7.8875 | -103.3593
7.8912 | -103.3477
7.8948 | -103.3381
7.8984 | -103.3276
7.902 | -103.318
7.9057 | -103.3072
7.9093 | -103.2963
7.9129 | -103.2817
7.9166 | -103.2689
7.9202 | -103.2605
7.9238 | -103.2505
7.9274 | -103.2351
7.9311 | -103.2161
7.9347 | -103.1973
7.9383 | -103.1711
7.942 | -103.1487
7.9456 | -103.126
7.9492 | -103.0995
7.9528 | -103.0696
7.9565 | -103.0458
7.9601 | -103.0291
7.9637 | -103.0175
7.9673 | -103.0098
7.971 | -103.0043
7.9746 | -103.0011
7.9782 | -103.0001
7.9819 | -103.0034
7.9855 | -103.0085
7.9891 | -103.0209
7.9927 | -103.0327
7.9964 | -103.0441
8 | -103.0583
8.0036 | -103.0734
8.0073 | -103.0899
8.0109 | -103.102
8.0145 | -103.1196
8.0181 | -103.1409
8.0218 | -103.1591
8.0254 | -103.1842
8.029 | -103.2138
8.0327 | -103.245
8.0363 | -103.2778
8.0399 | -103.3122
8.0435 | -103.3487
8.0472 | -103.3885
8.0508 | -103.4257
8.0544 | -103.4651
8.058 | -103.5077
8.0617 | -103.5452
8.0653 | -103.5769
8.0689 | -103.604
8.0726 | -103.6241
8.0762 | -103.6299
8.0798 | -103.6294
8.0834 | -103.6215
8.0871 | -103.6106
8.0907 | -103.5933
8.0943 | -103.5709
8.098 | -103.5437
8.1016 | -103.5102
8.1052 | -103.4739
8.1088 | -103.44
8.1125 | -103.4112
8.1161 | -103.3815
8.1197 | -103.3583
8.1234 | -103.3387
8.127 | -103.3196
8.1306 | -103.3062
8.1342 | -103.294
8.1379 | -103.2859
8.1415 | -103.2854
8.1451 | -103.2868
8.1488 | -103.2872
8.1524 | -103.2895
8.156 | -103.2899
8.1596 | -103.2901
8.1633 | -103.2921
8.1669 | -103.2942
8.1705 | -103.2983
8.1741 | -103.2949
8.1778 | -103.2909
8.1814 | -103.2914
8.185 | -103.2979
8.1887 | -103.3085
8.1923 | -103.3299
8.1959 | -103.3581
8.1995 | -103.3841
8.2032 | -103.4135
8.2068 | -103.444
8.2104 | -103.4818
8.2141 | -103.5234
8.2177 | -103.5674
8.2213 | -103.6139
8.2249 | -103.6555
8.2286 | -103.6864
8.2322 | -103.7119
8.2358 | -103.7349
8.2395 | -103.7529
8.2431 | -103.7635
8.2467 | -103.7682
8.2503 | -103.7616
8.254 | -103.7477
8.2576 | -103.7284
8.2612 | -103.7055
8.2649 | -103.6807
8.2685 | -103.6579
8.2721 | -103.6367
8.2757 | -103.6181
8.2794 | -103.5984
8.283 | -103.5792
8.2866 | -103.5661
8.2902 | -103.5532
8.2939 | -103.5454
8.2975 | -103.5363
8.3011 | -103.5227
8.3048 | -103.5087
8.3084 | -103.492
8.312 | -103.4639
8.3156 | -103.4358
8.3193 | -103.4058
8.3229 | -103.3772
8.3265 | -103.3472
8.3302 | -103.3088
8.3338 | -103.2732
8.3374 | -103.2395
8.341 | -103.2105
8.3447 | -103.1886
8.3483 | -103.1709
8.3519 | -103.1525
8.3556 | -103.138
8.3592 | -103.1243
8.3628 | -103.1178
8.3664 | -103.1176
8.3701 | -103.1219
8.3737 | -103.1381
8.3773 | -103.1476
8.381 | -103.1632
8.3846 | -103.1787
8.3882 | -103.195
8.3918 | -103.2098
8.3955 | -103.2235
8.3991 | -103.2379
8.4027 | -103.2553
8.4063 | -103.2742
8.41 | -103.2895
8.4136 | -103.3028
8.4172 | -103.3141
8.4209 | -103.3263
8.4245 | -103.3401
8.4281 | -103.3481
8.4317 | -103.3561
8.4354 | -103.3681
8.439 | -103.37
8.4426 | -103.3684
8.4463 | -103.3658
8.4499 | -103.3642
8.4535 | -103.3626
8.4571 | -103.3556
8.4608 | -103.3486
8.4644 | -103.3438
8.468 | -103.334
8.4717 | -103.3191
8.4753 | -103.306
8.4789 | -103.2892
8.4825 | -103.2715
8.4862 | -103.2485
8.4898 | -103.2213
8.4934 | -103.1987
8.4971 | -103.1799
8.5007 | -103.161
8.5043 | -103.1449
8.5079 | -103.1282
8.5116 | -103.1189
8.5152 | -103.1165
8.5188 | -103.1182
8.5224 | -103.1224
8.5261 | -103.1304
8.5297 | -103.1419
8.5333 | -103.152
8.537 | -103.163
8.5406 | -103.182
8.5442 | -103.2037
8.5478 | -103.2234
8.5515 | -103.2462
8.5551 | -103.2708
8.5587 | -103.294
8.5624 | -103.3179
8.566 | -103.3339
8.5696 | -103.3481
8.5732 | -103.356
8.5769 | -103.3532
8.5805 | -103.349
8.5841 | -103.3419
8.5878 | -103.3247
8.5914 | -103.3057
8.595 | -103.2886
8.5986 | -103.2763
8.6023 | -103.2636
8.6059 | -103.2514
8.6095 | -103.2407
8.6132 | -103.2386
8.6168 | -103.2415
8.6204 | -103.2444
8.624 | -103.2517
8.6277 | -103.2636
8.6313 | -103.2799
8.6349 | -103.3012
8.6385 | -103.3211
8.6422 | -103.3417
8.6458 | -103.3669
8.6494 | -103.3945
8.6531 | -103.4178
8.6567 | -103.4407
8.6603 | -103.4652
8.6639 | -103.4882
8.6676 | -103.5061
8.6712 | -103.5227
8.6748 | -103.5345
8.6785 | -103.5394
8.6821 | -103.5398
8.6857 | -103.5378
8.6893 | -103.5324
8.693 | -103.5329
8.6966 | -103.5324
8.7002 | -103.5273
8.7039 | -103.5241
8.7075 | -103.5195
8.7111 | -103.513
8.7147 | -103.5098
8.7184 | -103.5073
8.722 | -103.5059
8.7256 | -103.5046
8.7293 | -103.5041
8.7329 | -103.509
8.7365 | -103.5141
8.7401 | -103.5183
8.7438 | -103.5244
8.7474 | -103.5296
8.751 | -103.5371
8.7546 | -103.5434
8.7583 | -103.5459
8.7619 | -103.5521
8.7655 | -103.5571
8.7692 | -103.5534
8.7728 | -103.5472
8.7764 | -103.5379
8.78 | -103.5294
8.7837 | -103.5186
8.7873 | -103.5084
8.7909 | -103.4996
8.7946 | -103.4893
8.7982 | -103.4813
8.8018 | -103.4767
8.8054 | -103.4741
8.8091 | -103.4728
8.8127 | -103.4698
8.8163 | -103.4654
8.82 | -103.4652
8.8236 | -103.4698
8.8272 | -103.4774
8.8308 | -103.4875
8.8345 | -103.4927
8.8381 | -103.5001
8.8417 | -103.5104
8.8454 | -103.5152
8.849 | -103.5183
8.8526 | -103.514
8.8562 | -103.5077
8.8599 | -103.5011
8.8635 | -103.4938
8.8671 | -103.4876
8.8707 | -103.4821
8.8744 | -103.4774
8.878 | -103.4728
8.8816 | -103.4657
8.8853 | -103.4636
8.8889 | -103.4619
8.8925 | -103.4646
8.8961 | -103.4702
8.8998 | -103.4753
8.9034 | -103.4806
8.907 | -103.488
8.9107 | -103.4926
8.9143 | -103.4928
8.9179 | -103.4881
8.9215 | -103.4836
8.9252 | -103.4785
8.9288 | -103.4717
8.9324 | -103.4647
8.9361 | -103.4568
8.9397 | -103.4466
8.9433 | -103.4413
8.9469 | -103.4356
8.9506 | -103.4293
8.9542 | -103.4208
8.9578 | -103.4136
8.9615 | -103.4075
8.9651 | -103.4089
8.9687 | -103.408
8.9723 | -103.405
8.976 | -103.4015
8.9796 | -103.4017
8.9832 | -103.4058
8.9868 | -103.4037
8.9905 | -103.3978
8.9941 | -103.3908
8.9977 | -103.3814
9.0014 | -103.3742
9.005 | -103.3696
9.0086 | -103.3635
9.0122 | -103.3609
9.0159 | -103.3562
9.0195 | -103.3526
9.0231 | -103.3506
9.0268 | -103.353
9.0304 | -103.3565
9.034 | -103.3587
9.0376 | -103.3631
9.0413 | -103.3682
9.0449 | -103.3772
9.0485 | -103.387
9.0522 | -103.4011
9.0558 | -103.4127
9.0594 | -103.4262
9.063 | -103.4417
9.0667 | -103.4569
9.0703 | -103.4685
9.0739 | -103.484
9.0776 | -103.4948
9.0812 | -103.5128
9.0848 | -103.5344
9.0884 | -103.5598
9.0921 | -103.5884
9.0957 | -103.6184
9.0993 | -103.6489
9.1029 | -103.6808
9.1066 | -103.7118
9.1102 | -103.7475
9.1138 | -103.784
9.1175 | -103.8218
9.1211 | -103.8618
9.1247 | -103.9051
9.1283 | -103.947
9.132 | -103.9881
9.1356 | -104.0272
9.1392 | -104.0663
9.1429 | -104.1029
9.1465 | -104.1361
9.1501 | -104.1666
9.1537 | -104.1907
9.1574 | -104.2097
9.161 | -104.2239
9.1646 | -104.2382
9.1683 | -104.2489
9.1719 | -104.2537
9.1755 | -104.2591
9.1791 | -104.2589
9.1828 | -104.2588
9.1864 | -104.2567
9.19 | -104.2531
9.1937 | -104.2466
9.1973 | -104.2387
9.2009 | -104.2326
9.2045 | -104.2307
9.2082 | -104.2281
9.2118 | -104.2235
9.2154 | -104.2161
9.219 | -104.2071
9.2227 | -104.1994
9.2263 | -104.1971
9.2299 | -104.1916
9.2336 | -104.1845
9.2372 | -104.1768
9.2408 | -104.1708
9.2444 | -104.1601
9.2481 | -104.147
9.2517 | -104.1347
9.2553 | -104.1248
9.259 | -104.1123
9.2626 | -104.1006
9.2662 | -104.0919
9.2698 | -104.0824
9.2735 | -104.0781
9.2771 | -104.0793
9.2807 | -104.0878
9.2844 | -104.0978
9.288 | -104.1074
9.2916 | -104.1229
9.2952 | -104.1363
9.2989 | -104.149
9.3025 | -104.1651
9.3061 | -104.1784
9.3098 | -104.1887
9.3134 | -104.1991
9.317 | -104.2034
9.3206 | -104.2033
9.3243 | -104.1947
9.3279 | -104.1792
9.3315 | -104.1635
9.3351 | -104.1402
9.3388 | -104.1085
9.3424 | -104.069
9.346 | -104.0267
9.3497 | -103.9762
9.3533 | -103.9236
9.3569 | -103.8698
9.3605 | -103.8156
9.3642 | -103.7665
9.3678 | -103.7195
9.3714 | -103.6768
9.3751 | -103.635
9.3787 | -103.6017
9.3823 | -103.5749
9.3859 | -103.5538
9.3896 | -103.5432
9.3932 | -103.5447
9.3968 | -103.5571
9.4005 | -103.5713
9.4041 | -103.5919
9.4077 | -103.6203
9.4113 | -103.6556
9.415 | -103.6934
9.4186 | -103.7426
9.4222 | -103.7852
9.4259 | -103.834
9.4295 | -103.8795
9.4331 | -103.9226
9.4367 | -103.9599
9.4404 | -103.9911
9.444 | -104.0195
9.4476 | -104.0408
9.4512 | -104.056
9.4549 | -104.0692
9.4585 | -104.0758
9.4621 | -104.0799
9.4658 | -104.0774
9.4694 | -104.0659
9.473 | -104.0549
9.4766 | -104.0432
9.4803 | -104.0314
9.4839 | -104.017
9.4875 | -104.0031
9.4912 | -103.9903
9.4948 | -103.9802
9.4984 | -103.9727
9.502 | -103.9662
9.5057 | -103.9629
9.5093 | -103.9645
9.5129 | -103.9711
9.5166 | -103.9757
9.5202 | -103.9824
9.5238 | -103.9927
9.5274 | -104.0067
9.5311 | -104.0198
9.5347 | -104.03
9.5383 | -104.0441
9.5419 | -104.0615
9.5456 | -104.077
9.5492 | -104.0917
9.5528 | -104.1051
9.5565 | -104.1194
9.5601 | -104.1342
9.5637 | -104.1499
9.5673 | -104.1645
9.571 | -104.1816
9.5746 | -104.2002
9.5782 | -104.2181
9.5819 | -104.2325
9.5855 | -104.2469
9.5891 | -104.2633
9.5927 | -104.2798
9.5964 | -104.2881
9.6 | -104.2974
9.6036 | -104.3045
9.6073 | -104.3077
9.6109 | -104.3091
9.6145 | -104.3064
9.6181 | -104.3028
9.6218 | -104.2924
9.6254 | -104.2758
9.629 | -104.2597
9.6327 | -104.2387
9.6363 | -104.2143
9.6399 | -104.1901
9.6435 | -104.1695
9.6472 | -104.1492
9.6508 | -104.125
9.6544 | -104.1049
9.658 | -104.0902
9.6617 | -104.0741
9.6653 | -104.0591
9.6689 | -104.0459
9.6726 | -104.039
9.6762 | -104.0414
9.6798 | -104.0441
9.6834 | -104.0474
9.6871 | -104.0523
9.6907 | -104.061
9.6943 | -104.0724
9.698 | -104.0816
9.7016 | -104.0915
9.7052 | -104.1042
9.7088 | -104.1194
9.7125 | -104.1348
9.7161 | -104.1515
9.7197 | -104.1655
9.7234 | -104.176
9.727 | -104.1819
9.7306 | -104.1816
9.7342 | -104.1772
9.7379 | -104.1748
9.7415 | -104.1679
9.7451 | -104.1599
9.7488 | -104.1515
9.7524 | -104.1393
9.756 | -104.1264
9.7596 | -104.112
9.7633 | -104.1024
9.7669 | -104.0969
9.7705 | -104.0908
9.7742 | -104.0878
9.7778 | -104.0857
9.7814 | -104.0809
9.785 | -104.0718
9.7887 | -104.0636
9.7923 | -104.0611
9.7959 | -104.0585
9.7995 | -104.0583
9.8032 | -104.0628
9.8068 | -104.0657
9.8104 | -104.0643
9.8141 | -104.0615
9.8177 | -104.0604
9.8213 | -104.0623
9.8249 | -104.0626
9.8286 | -104.0641
9.8322 | -104.0601
9.8358 | -104.058
9.8395 | -104.0579
9.8431 | -104.0585
9.8467 | -104.0565
9.8503 | -104.0553
9.854 | -104.0556
9.8576 | -104.0587
9.8612 | -104.0677
9.8649 | -104.0794
9.8685 | -104.0886
9.8721 | -104.1042
9.8757 | -104.1236
9.8794 | -104.1448
9.883 | -104.1677
9.8866 | -104.1955
9.8902 | -104.2269
9.8939 | -104.2601
9.8975 | -104.2949
9.9011 | -104.3332
9.9048 | -104.3682
9.9084 | -104.4039
9.912 | -104.4343
9.9156 | -104.4618
9.9193 | -104.4879
9.9229 | -104.5145
9.9265 | -104.5417
9.9302 | -104.5597
9.9338 | -104.5717
9.9374 | -104.579
9.941 | -104.5805
9.9447 | -104.5659
9.9483 | -104.557
9.9519 | -104.5494
9.9556 | -104.5444
9.9592 | -104.5365
9.9628 | -104.5224
9.9664 | -104.5085
9.9701 | -104.4976
9.9737 | -104.4855
9.9773 | -104.4814
9.981 | -104.4781
9.9846 | -104.476
9.9882 | -104.4726
9.9918 | -104.4725
9.9955 | -104.4754
9.9991 | -104.4802
10.0027 | -104.4847
10.0063 | -104.4877
10.01 | -104.4961
10.0136 | -104.5123
10.0172 | -104.5257
10.0209 | -104.5382
10.0245 | -104.5487
10.0281 | -104.5614
10.0317 | -104.5777
10.0354 | -104.5961
10.039 | -104.6141
10.0426 | -104.6384
10.0463 | -104.6613
10.0499 | -104.6883
10.0535 | -104.7175
10.0571 | -104.7438
10.0608 | -104.7712
10.0644 | -104.7987
10.068 | -104.8244
10.0717 | -104.8435
10.0753 | -104.8655
10.0789 | -104.8854
10.0825 | -104.9004
10.0862 | -104.9119
10.0898 | -104.9207
10.0934 | -104.9286
10.0971 | -104.9331
10.1007 | -104.9311
10.1043 | -104.9246
10.1079 | -104.9149
10.1116 | -104.904
10.1152 | -104.887
10.1188 | -104.8683
10.1224 | -104.8544
10.1261 | -104.8371
10.1297 | -104.8149
10.1333 | -104.7909
10.137 | -104.7734
10.1406 | -104.7576
10.1442 | -104.745
10.1478 | -104.7337
10.1515 | -104.7242
10.1551 | -104.7176
10.1587 | -104.7117
10.1624 | -104.6999
10.166 | -104.6939
10.1696 | -104.6916
10.1732 | -104.6922
10.1769 | -104.6899
10.1805 | -104.6844
10.1841 | -104.6741
10.1878 | -104.6636
10.1914 | -104.644
10.195 | -104.6256
10.1986 | -104.6062
10.2023 | -104.5865
10.2059 | -104.569
10.2095 | -104.5438
10.2132 | -104.5143
10.2168 | -104.4846
10.2204 | -104.4613
10.224 | -104.4375
10.2277 | -104.4108
10.2313 | -104.3869
10.2349 | -104.3697
10.2385 | -104.3588
10.2422 | -104.349
10.2458 | -104.3442
10.2494 | -104.3458
10.2531 | -104.3529
10.2567 | -104.3643
10.2603 | -104.3824
10.2639 | -104.4049
10.2676 | -104.4289
10.2712 | -104.4531
10.2748 | -104.4824
10.2785 | -104.5182
10.2821 | -104.5539
10.2857 | -104.5868
10.2893 | -104.6218
10.293 | -104.6565
10.2966 | -104.692
10.3002 | -104.7221
10.3039 | -104.7499
10.3075 | -104.7782
10.3111 | -104.8051
10.3147 | -104.8232
10.3184 | -104.841
10.322 | -104.8544
10.3256 | -104.8658
10.3293 | -104.8746
10.3329 | -104.8826
10.3365 | -104.8861
10.3401 | -104.8947
10.3438 | -104.8948
10.3474 | -104.8904
10.351 | -104.8843
10.3546 | -104.8787
10.3583 | -104.876
10.3619 | -104.8731
10.3655 | -104.8716
10.3692 | -104.8658
10.3728 | -104.86
10.3764 | -104.8565
10.38 | -104.8519
10.3837 | -104.8394
10.3873 | -104.8254
10.3909 | -104.8142
10.3946 | -104.7991
10.3982 | -104.7789
10.4018 | -104.7602
10.4054 | -104.7433
10.4091 | -104.7234
10.4127 | -104.7042
10.4163 | -104.6849
10.42 | -104.6655
10.4236 | -104.6475
10.4272 | -104.6299
10.4308 | -104.6136
10.4345 | -104.5993
10.4381 | -104.5853
10.4417 | -104.5715
10.4454 | -104.5607
10.449 | -104.5606
10.4526 | -104.5599
10.4562 | -104.5552
10.4599 | -104.5574
10.4635 | -104.5666
10.4671 | -104.5844
10.4707 | -104.5933
10.4744 | -104.6055
10.478 | -104.6182
10.4816 | -104.6291
10.4853 | -104.6439
10.4889 | -104.6608
10.4925 | -104.6738
10.4961 | -104.69
10.4998 | -104.7032
10.5034 | -104.7089
10.507 | -104.7128
10.5107 | -104.7153
10.5143 | -104.7117
10.5179 | -104.7038
10.5215 | -104.6967
10.5252 | -104.685
10.5288 | -104.6707
10.5324 | -104.6543
10.5361 | -104.6395
10.5397 | -104.6198
10.5433 | -104.5988
10.5469 | -104.5796
10.5506 | -104.5641
10.5542 | -104.5459
10.5578 | -104.5286
10.5615 | -104.5145
10.5651 | -104.5004
10.5687 | -104.4878
10.5723 | -104.4784
10.576 | -104.4715
10.5796 | -104.4704
10.5832 | -104.4723
10.5868 | -104.4731
10.5905 | -104.4719
10.5941 | -104.4675
10.5977 | -104.464
10.6014 | -104.4583
10.605 | -104.4543
10.6086 | -104.4485
10.6122 | -104.4414
10.6159 | -104.431
10.6195 | -104.4234
10.6231 | -104.4189
10.6268 | -104.4123
10.6304 | -104.4074
10.634 | -104.4015
10.6376 | -104.3983
10.6413 | -104.398
10.6449 | -104.3993
10.6485 | -104.4026
10.6522 | -104.4071
10.6558 | -104.4159
10.6594 | -104.428
10.663 | -104.44
10.6667 | -104.4537
10.6703 | -104.4729
10.6739 | -104.489
10.6776 | -104.5035
10.6812 | -104.5213
10.6848 | -104.5392
10.6884 | -104.5541
10.6921 | -104.5639
10.6957 | -104.5737
10.6993 | -104.582
10.7029 | -104.5871
10.7066 | -104.5898
10.7102 | -104.5905
10.7138 | -104.5885
10.7175 | -104.5815
10.7211 | -104.5763
10.7247 | -104.5689
10.7283 | -104.5592
10.732 | -104.5479
10.7356 | -104.5328
10.7392 | -104.5197
10.7429 | -104.5074
10.7465 | -104.4992
10.7501 | -104.4946
10.7537 | -104.4886
10.7574 | -104.4821
10.761 | -104.4797
10.7646 | -104.4836
10.7683 | -104.489
10.7719 | -104.4929
10.7755 | -104.4953
10.7791 | -104.5005
10.7828 | -104.5115
10.7864 | -104.5261
10.79 | -104.5492
10.7937 | -104.5686
10.7973 | -104.5868
10.8009 | -104.6075
10.8045 | -104.6305
10.8082 | -104.6539
10.8118 | -104.674
10.8154 | -104.6962
10.819 | -104.7174
10.8227 | -104.7311
10.8263 | -104.7371
10.8299 | -104.7485
10.8336 | -104.7581
10.8372 | -104.7634
10.8408 | -104.7646
10.8444 | -104.7624
10.8481 | -104.7621
10.8517 | -104.7549
10.8553 | -104.7444
10.859 | -104.7346
10.8626 | -104.7198
10.8662 | -104.704
10.8698 | -104.6888
10.8735 | -104.6697
10.8771 | -104.646
10.8807 | -104.6227
10.8844 | -104.6013
10.888 | -104.5824
10.8916 | -104.5676
10.8952 | -104.5534
10.8989 | -104.5416
10.9025 | -104.5327
10.9061 | -104.5295
10.9098 | -104.5304
10.9134 | -104.5337
10.917 | -104.5407
10.9206 | -104.5486
10.9243 | -104.5588
10.9279 | -104.5763
10.9315 | -104.5967
10.9351 | -104.6148
10.9388 | -104.6296
10.9424 | -104.6468
10.946 | -104.6616
10.9497 | -104.6779
10.9533 | -104.6901
10.9569 | -104.6995
10.9605 | -104.7074
10.9642 | -104.7138
10.9678 | -104.718
10.9714 | -104.7171
10.9751 | -104.7173
10.9787 | -104.72
10.9823 | -104.7213
10.9859 | -104.716
10.9896 | -104.6998
10.9932 | -104.6894
10.9968 | -104.68
11.0005 | -104.674
11.0041 | -104.667
11.0077 | -104.6567
11.0113 | -104.6474
11.015 | -104.6368
11.0186 | -104.6228
11.0222 | -104.61
11.0259 | -104.5986
11.0295 | -104.5867
11.0331 | -104.5787
11.0367 | -104.5699
11.0404 | -104.5612
11.044 | -104.5526
11.0476 | -104.5416
11.0512 | -104.5332
11.0549 | -104.5306
11.0585 | -104.5272
11.0621 | -104.5263
11.0658 | -104.5264
11.0694 | -104.526
11.073 | -104.5287
11.0766 | -104.5335
11.0803 | -104.5409
11.0839 | -104.5508
11.0875 | -104.56
11.0912 | -104.567
11.0948 | -104.5724
11.0984 | -104.5819
11.102 | -104.5913
11.1057 | -104.601
11.1093 | -104.6106
11.1129 | -104.6206
11.1166 | -104.6341
11.1202 | -104.6501
11.1238 | -104.6683
11.1274 | -104.6877
11.1311 | -104.7074
11.1347 | -104.7259
11.1383 | -104.746
11.1419 | -104.7687
11.1456 | -104.7916
11.1492 | -104.8188
11.1528 | -104.847
11.1565 | -104.8729
11.1601 | -104.8953
11.1637 | -104.9133
11.1673 | -104.933
11.171 | -104.9488
11.1746 | -104.9599
11.1782 | -104.9719
11.1819 | -104.9825
11.1855 | -104.9927
11.1891 | -104.9956
11.1927 | -104.9966
11.1964 | -104.9955
11.2 | -104.9894
11.2036 | -104.9786
11.2073 | -104.9645
11.2109 | -104.95
11.2145 | -104.933
11.2181 | -104.9123
11.2218 | -104.8943
11.2254 | -104.8721
11.229 | -104.8478
11.2327 | -104.8245
11.2363 | -104.8022
11.2399 | -104.7795
11.2435 | -104.7614
11.2472 | -104.744
11.2508 | -104.7304
11.2544 | -104.72
11.2581 | -104.7137
11.2617 | -104.7061
11.2653 | -104.7002
11.2689 | -104.6967
11.2726 | -104.7004
11.2762 | -104.7028
11.2798 | -104.7011
11.2834 | -104.699
11.2871 | -104.7014
11.2907 | -104.704
11.2943 | -104.7054
11.298 | -104.7061
11.3016 | -104.7077
11.3052 | -104.71
11.3088 | -104.7097
11.3125 | -104.7051
11.3161 | -104.7022
11.3197 | -104.6995
11.3234 | -104.6932
11.327 | -104.69
11.3306 | -104.6871
11.3342 | -104.6829
11.3379 | -104.6797
11.3415 | -104.671
11.3451 | -104.6644
11.3488 | -104.6592
11.3524 | -104.6549
11.356 | -104.6506
11.3596 | -104.647
11.3633 | -104.6448
11.3669 | -104.6416
11.3705 | -104.6404
11.3742 | -104.6412
11.3778 | -104.638
11.3814 | -104.6381
11.385 | -104.6413
11.3887 | -104.6481
11.3923 | -104.6554
11.3959 | -104.6625
11.3995 | -104.6686
11.4032 | -104.6737
11.4068 | -104.6783
11.4104 | -104.6813
11.4141 | -104.6853
11.4177 | -104.6896
11.4213 | -104.6941
11.4249 | -104.7001
11.4286 | -104.7033
11.4322 | -104.7049
11.4358 | -104.7062
11.4395 | -104.7082
11.4431 | -104.713
11.4467 | -104.7147
11.4503 | -104.717
11.454 | -104.7217
11.4576 | -104.7243
11.4612 | -104.7257
11.4649 | -104.7278
11.4685 | -104.7298
11.4721 | -104.7325
11.4757 | -104.7347
11.4794 | -104.7368
11.483 | -104.7368
11.4866 | -104.7392
11.4902 | -104.738
11.4939 | -104.7369
11.4975 | -104.7345
11.5011 | -104.7331
11.5048 | -104.7301
11.5084 | -104.7262
11.512 | -104.7227
11.5156 | -104.7159
11.5193 | -104.7128
11.5229 | -104.711
11.5265 | -104.7091
11.5302 | -104.7097
11.5338 | -104.71
11.5374 | -104.7096
11.541 | -104.709
11.5447 | -104.7129
11.5483 | -104.7165
11.5519 | -104.7187
11.5556 | -104.7211
11.5592 | -104.7262
11.5628 | -104.7326
11.5664 | -104.7408
11.5701 | -104.7494
11.5737 | -104.7551
11.5773 | -104.7646
11.581 | -104.7734
11.5846 | -104.7803
11.5882 | -104.7878
11.5918 | -104.7919
11.5955 | -104.7954
11.5991 | -104.8014
11.6027 | -104.8063
11.6063 | -104.8116
11.61 | -104.8134
11.6136 | -104.819
11.6172 | -104.8237
11.6209 | -104.8302
11.6245 | -104.8366
11.6281 | -104.8441
11.6317 | -104.8495
11.6354 | -104.8547
11.639 | -104.8599
11.6426 | -104.8676
11.6463 | -104.8726
11.6499 | -104.8787
11.6535 | -104.8839
11.6571 | -104.8888
11.6608 | -104.8929
11.6644 | -104.8985
11.668 | -104.902
11.6717 | -104.9045
11.6753 | -104.9038
11.6789 | -104.9049
11.6825 | -104.9023
11.6862 | -104.9001
11.6898 | -104.8969
11.6934 | -104.8912
11.6971 | -104.8863
11.7007 | -104.878
11.7043 | -104.8699
11.7079 | -104.8572
11.7116 | -104.8434
11.7152 | -104.8307
11.7188 | -104.8165
11.7224 | -104.8011
11.7261 | -104.7849
11.7297 | -104.7697
11.7333 | -104.7518
11.737 | -104.732
11.7406 | -104.711
11.7442 | -104.6928
11.7478 | -104.6739
11.7515 | -104.6596
11.7551 | -104.6508
11.7587 | -104.6397
11.7624 | -104.6338
11.766 | -104.6349
11.7696 | -104.6397
11.7732 | -104.6458
11.7769 | -104.655
11.7805 | -104.6695
11.7841 | -104.6834
11.7878 | -104.7011
11.7914 | -104.7223
11.795 | -104.752
11.7986 | -104.783
11.8023 | -104.814
11.8059 | -104.8484
11.8095 | -104.883
11.8132 | -104.9197
11.8168 | -104.9546
11.8204 | -104.9896
11.824 | -105.0233
11.8277 | -105.0513
11.8313 | -105.0772
11.8349 | -105.1004
11.8385 | -105.1208
11.8422 | -105.1397
11.8458 | -105.1545
11.8494 | -105.1622
11.8531 | -105.169
11.8567 | -105.1697
11.8603 | -105.1652
11.8639 | -105.1589
11.8676 | -105.147
11.8712 | -105.1353
11.8748 | -105.1192
11.8785 | -105.1003
11.8821 | -105.0792
11.8857 | -105.0559
11.8893 | -105.0313
11.893 | -105.0058
11.8966 | -104.9762
11.9002 | -104.9462
11.9039 | -104.9241
11.9075 | -104.9017
11.9111 | -104.8823
11.9147 | -104.8653
11.9184 | -104.8466
11.922 | -104.8238
11.9256 | -104.8018
11.9293 | -104.7868
11.9329 | -104.7735
11.9365 | -104.7601
11.9401 | -104.7465
11.9438 | -104.7347
11.9474 | -104.7261
11.951 | -104.7162
11.9546 | -104.7071
11.9583 | -104.7002
11.9619 | -104.6958
11.9655 | -104.6888
11.9692 | -104.6867
11.9728 | -104.6894
11.9764 | -104.6925
11.98 | -104.6997
11.9837 | -104.7078
11.9873 | -104.7196
11.9909 | -104.7347
11.9946 | -104.7545
11.9982 | -104.7791
12.0018 | -104.8025
12.0054 | -104.827
12.0091 | -104.8521
12.0127 | -104.8774
12.0163 | -104.9028
12.02 | -104.9252
12.0236 | -104.9448
12.0272 | -104.9655
12.0308 | -104.9846
12.0345 | -105.0014
12.0381 | -105.0041
12.0417 | -105.0152
12.0454 | -105.027
12.049 | -105.0381
12.0526 | -105.0472
12.0562 | -105.0539
12.0599 | -105.0547
12.0635 | -105.0544
12.0671 | -105.0548
12.0707 | -105.0583
12.0744 | -105.0556
12.078 | -105.0492
12.0816 | -105.0405
12.0853 | -105.0306
12.0889 | -105.0205
12.0925 | -105.0095
12.0961 | -105.0013
12.0998 | -104.9936
12.1034 | -104.984
12.107 | -104.9728
12.1107 | -104.9619
12.1143 | -104.9517
12.1179 | -104.9403
12.1215 | -104.9325
12.1252 | -104.9257
12.1288 | -104.9193
12.1324 | -104.9084
12.1361 | -104.8947
12.1397 | -104.8848
12.1433 | -104.8776
12.1469 | -104.8677
12.1506 | -104.8587
12.1542 | -104.8515
12.1578 | -104.8441
12.1615 | -104.8376
12.1651 | -104.8318
12.1687 | -104.8268
12.1723 | -104.8239
12.176 | -104.8236
12.1796 | -104.8253
12.1832 | -104.8261
12.1868 | -104.8234
12.1905 | -104.8237
12.1941 | -104.8274
12.1977 | -104.8286
12.2014 | -104.8274
12.205 | -104.8258
12.2086 | -104.8229
12.2122 | -104.8191
12.2159 | -104.8193
12.2195 | -104.8196
12.2231 | -104.8145
12.2268 | -104.8075
12.2304 | -104.8043
12.234 | -104.8001
12.2376 | -104.8005
12.2413 | -104.8028
12.2449 | -104.8049
12.2485 | -104.8069
12.2522 | -104.812
12.2558 | -104.815
12.2594 | -104.8224
12.263 | -104.832
12.2667 | -104.8412
12.2703 | -104.8548
12.2739 | -104.8691
12.2776 | -104.8831
12.2812 | -104.8932
12.2848 | -104.905
12.2884 | -104.9224
12.2921 | -104.9403
12.2957 | -104.96
12.2993 | -104.9794
12.3029 | -105.0003
12.3066 | -105.021
12.3102 | -105.0423
12.3138 | -105.0631
12.3175 | -105.0825
12.3211 | -105.0968
12.3247 | -105.1103
12.3283 | -105.1217
12.332 | -105.1298
12.3356 | -105.1333
12.3392 | -105.1359
12.3429 | -105.137
12.3465 | -105.1346
12.3501 | -105.128
12.3537 | -105.1172
12.3574 | -105.1039
12.361 | -105.0881
12.3646 | -105.0752
12.3683 | -105.0605
12.3719 | -105.0444
12.3755 | -105.0312
12.3791 | -105.0143
12.3828 | -104.9966
12.3864 | -104.9825
12.39 | -104.9698
12.3937 | -104.9542
12.3973 | -104.9388
12.4009 | -104.9296
12.4045 | -104.9235
12.4082 | -104.9136
12.4118 | -104.9061
12.4154 | -104.8973
12.419 | -104.8886
12.4227 | -104.8826
12.4263 | -104.8787
12.4299 | -104.8788
12.4336 | -104.8807
12.4372 | -104.8845
12.4408 | -104.8871
12.4444 | -104.8886
12.4481 | -104.8927
12.4517 | -104.8952
12.4553 | -104.9017
12.459 | -104.9127
12.4626 | -104.9227
12.4662 | -104.9297
12.4698 | -104.9362
12.4735 | -104.9391
12.4771 | -104.9386
12.4807 | -104.9395
12.4844 | -104.9387
12.488 | -104.9353
12.4916 | -104.9307
12.4952 | -104.9238
12.4989 | -104.9147
12.5025 | -104.9074
12.5061 | -104.9004
12.5098 | -104.8903
12.5134 | -104.8802
12.517 | -104.872
12.5206 | -104.8619
12.5243 | -104.852
12.5279 | -104.8454
12.5315 | -104.8434
12.5351 | -104.8417
12.5388 | -104.8426
12.5424 | -104.8458
12.546 | -104.8495
12.5497 | -104.8555
12.5533 | -104.8648
12.5569 | -104.8793
12.5605 | -104.9017
12.5642 | -104.9167
12.5678 | -104.9325
12.5714 | -104.9512
12.5751 | -104.971
12.5787 | -104.9922
12.5823 | -105.0145
12.5859 | -105.0373
12.5896 | -105.0609
12.5932 | -105.0829
12.5968 | -105.1059
12.6005 | -105.1302
12.6041 | -105.154
12.6077 | -105.173
12.6113 | -105.1911
12.615 | -105.2109
12.6186 | -105.2295
12.6222 | -105.2494
12.6258 | -105.268
12.6295 | -105.2806
12.6331 | -105.2912
12.6367 | -105.297
12.6404 | -105.3038
12.644 | -105.3069
12.6476 | -105.3097
12.6512 | -105.3108
12.6549 | -105.3104
12.6585 | -105.3063
12.6621 | -105.2982
12.6658 | -105.2898
12.6694 | -105.2795
12.673 | -105.269
12.6766 | -105.2574
12.6803 | -105.2447
12.6839 | -105.2297
12.6875 | -105.2143
12.6912 | -105.2023
12.6948 | -105.1919
12.6984 | -105.1801
12.702 | -105.1733
12.7057 | -105.1671
12.7093 | -105.1623
12.7129 | -105.1587
12.7166 | -105.1577
12.7202 | -105.1575
12.7238 | -105.1584
12.7274 | -105.1592
12.7311 | -105.1583
12.7347 | -105.1591
12.7383 | -105.1629
12.7419 | -105.1687
12.7456 | -105.1732
12.7492 | -105.1784
12.7528 | -105.1833
12.7565 | -105.1879
12.7601 | -105.1946
12.7637 | -105.1985
12.7673 | -105.1994
12.771 | -105.2029
12.7746 | -105.2069
12.7782 | -105.2105
12.7819 | -105.2091
12.7855 | -105.2034
12.7891 | -105.1955
12.7927 | -105.1896
12.7964 | -105.1824
12.8 | -105.173
12.8036 | -105.16
12.8073 | -105.1448
12.8109 | -105.1317
12.8145 | -105.1201
12.8181 | -105.1078
12.8218 | -105.0933
12.8254 | -105.0755
12.829 | -105.0582
12.8327 | -105.0378
12.8363 | -105.0192
12.8399 | -105.0056
12.8435 | -104.9943
12.8472 | -104.9805
12.8508 | -104.9704
12.8544 | -104.9638
12.8581 | -104.9551
12.8617 | -104.9479
12.8653 | -104.9427
12.8689 | -104.9429
12.8726 | -104.9436
12.8762 | -104.9453
12.8798 | -104.9494
12.8834 | -104.9546
12.8871 | -104.959
12.8907 | -104.9643
12.8943 | -104.9705
12.898 | -104.9745
12.9016 | -104.9795
12.9052 | -104.9851
12.9088 | -104.9898
12.9125 | -104.9971
12.9161 | -105.0083
12.9197 | -105.0136
12.9234 | -105.015
12.927 | -105.0167
12.9306 | -105.0179
12.9342 | -105.0184
12.9379 | -105.0148
12.9415 | -105.0088
12.9451 | -105.0026
12.9488 | -104.9933
12.9524 | -104.9822
12.956 | -104.971
12.9596 | -104.9593
12.9633 | -104.9445
12.9669 | -104.9313
12.9705 | -104.9205
12.9741 | -104.9069
12.9778 | -104.8936
12.9814 | -104.8799
12.985 | -104.8693
12.9887 | -104.8623
12.9923 | -104.857
12.9959 | -104.849
12.9995 | -104.8412
13.0032 | -104.8343
13.0068 | -104.8321
13.0104 | -104.8306
13.0141 | -104.8341
13.0177 | -104.8384
13.0213 | -104.842
13.0249 | -104.8448
13.0286 | -104.8482
13.0322 | -104.8552
13.0358 | -104.8597
13.0395 | -104.8651
13.0431 | -104.8727
13.0467 | -104.8798
13.0503 | -104.888
13.054 | -104.8909
13.0576 | -104.8934
13.0612 | -104.9002
13.0649 | -104.9041
13.0685 | -104.906
13.0721 | -104.9097
13.0757 | -104.914
13.0794 | -104.9157
13.083 | -104.9194
13.0866 | -104.9255
13.0902 | -104.9311
13.0939 | -104.9355
13.0975 | -104.9381
13.1011 | -104.9374
13.1048 | -104.9355
13.1084 | -104.9371
13.112 | -104.938
13.1156 | -104.9383
13.1193 | -104.933
13.1229 | -104.928
13.1265 | -104.9264
13.1302 | -104.9277
13.1338 | -104.9247
13.1374 | -104.9219
13.141 | -104.9202
13.1447 | -104.9155
13.1483 | -104.9085
13.1519 | -104.9016
13.1556 | -104.8933
13.1592 | -104.8824
13.1628 | -104.8748
13.1664 | -104.8651
13.1701 | -104.8523
13.1737 | -104.84
13.1773 | -104.8294
13.181 | -104.8203
13.1846 | -104.8142
13.1882 | -104.809
13.1918 | -104.8068
13.1955 | -104.8024
13.1991 | -104.8001
13.2027 | -104.8003
13.2063 | -104.8033
13.21 | -104.808
13.2136 | -104.8155
13.2172 | -104.8269
13.2209 | -104.8394
13.2245 | -104.851
13.2281 | -104.8605
13.2317 | -104.8711
13.2354 | -104.8879
13.239 | -104.9058
13.2426 | -104.926
13.2463 | -104.9457
13.2499 | -104.9635
13.2535 | -104.9804
13.2571 | -104.9985
13.2608 | -105.0108
13.2644 | -105.0234
13.268 | -105.0349
13.2717 | -105.0402
13.2753 | -105.0443
13.2789 | -105.0469
13.2825 | -105.0497
13.2862 | -105.0508
13.2898 | -105.0502
13.2934 | -105.0487
13.2971 | -105.046
13.3007 | -105.0414
13.3043 | -105.0352
13.3079 | -105.0287
13.3116 | -105.0259
13.3152 | -105.0209
13.3188 | -105.0128
13.3224 | -105.0034
13.3261 | -104.9972
13.3297 | -104.9944
13.3333 | -104.9931
13.337 | -104.9924
13.3406 | -104.9934
13.3442 | -104.997
13.3478 | -105.0026
13.3515 | -105.0078
13.3551 | -105.0136
13.3587 | -105.0208
13.3624 | -105.0329
13.366 | -105.0429
13.3696 | -105.0544
13.3732 | -105.0671
13.3769 | -105.0813
13.3805 | -105.0989
13.3841 | -105.1161
13.3878 | -105.135
13.3914 | -105.1536
13.395 | -105.1707
13.3986 | -105.1876
13.4023 | -105.2074
13.4059 | -105.2244
13.4095 | -105.2385
13.4132 | -105.255
13.4168 | -105.2717
13.4204 | -105.2853
13.424 | -105.2967
13.4277 | -105.2997
13.4313 | -105.2998
13.4349 | -105.2973
13.4385 | -105.2909
13.4422 | -105.2795
13.4458 | -105.2696
13.4494 | -105.2577
13.4531 | -105.2486
13.4567 | -105.2318
13.4603 | -105.214
13.4639 | -105.1964
13.4676 | -105.1766
13.4712 | -105.1568
13.4748 | -105.1335
13.4785 | -105.111
13.4821 | -105.0903
13.4857 | -105.0704
13.4893 | -105.0505
13.493 | -105.0335
13.4966 | -105.0211
13.5002 | -105.0083
13.5039 | -104.9952
13.5075 | -104.9837
13.5111 | -104.9748
13.5147 | -104.968
13.5184 | -104.9621
13.522 | -104.9595
13.5256 | -104.959
13.5293 | -104.9601
13.5329 | -104.9643
13.5365 | -104.9688
13.5401 | -104.975
13.5438 | -104.984
13.5474 | -104.989
13.551 | -104.9945
13.5546 | -105.0016
13.5583 | -105.0128
13.5619 | -105.0214
13.5655 | -105.0319
13.5692 | -105.0431
13.5728 | -105.0533
13.5764 | -105.0643
13.58 | -105.0773
13.5837 | -105.0893
13.5873 | -105.102
13.5909 | -105.1171
13.5946 | -105.1321
13.5982 | -105.1471
13.6018 | -105.1629
13.6054 | -105.1823
13.6091 | -105.196
13.6127 | -105.2142
13.6163 | -105.2339
13.62 | -105.2519
13.6236 | -105.2716
13.6272 | -105.2914
13.6308 | -105.3116
13.6345 | -105.329
13.6381 | -105.3456
13.6417 | -105.3617
13.6454 | -105.3757
13.649 | -105.3879
13.6526 | -105.4032
13.6562 | -105.4167
13.6599 | -105.4294
13.6635 | -105.4405
13.6671 | -105.4521
13.6707 | -105.4631
13.6744 | -105.473
13.678 | -105.4813
13.6816 | -105.4883
13.6853 | -105.4938
13.6889 | -105.4953
13.6925 | -105.4931
13.6961 | -105.4911
13.6998 | -105.4904
13.7034 | -105.4867
13.707 | -105.4795
13.7107 | -105.4713
13.7143 | -105.4657
13.7179 | -105.458
13.7215 | -105.4496
13.7252 | -105.4436
13.7288 | -105.4372
13.7324 | -105.4308
13.7361 | -105.4255
13.7397 | -105.4205
13.7433 | -105.4172
13.7469 | -105.4151
13.7506 | -105.4117
13.7542 | -105.4069
13.7578 | -105.403
13.7615 | -105.398
13.7651 | -105.3944
13.7687 | -105.3916
13.7723 | -105.3884
13.776 | -105.3855
13.7796 | -105.3876
13.7832 | -105.3884
13.7868 | -105.3901
13.7905 | -105.3914
13.7941 | -105.3948
13.7977 | -105.3987
13.8014 | -105.4038
13.805 | -105.4123
13.8086 | -105.4215
13.8122 | -105.4262
13.8159 | -105.433
13.8195 | -105.4398
13.8231 | -105.444
13.8268 | -105.4485
13.8304 | -105.4572
13.834 | -105.4665
13.8376 | -105.4739
13.8413 | -105.4784
13.8449 | -105.4834
13.8485 | -105.488
13.8522 | -105.4903
13.8558 | -105.4863
13.8594 | -105.4815
13.863 | -105.4781
13.8667 | -105.4735
13.8703 | -105.4701
13.8739 | -105.4642
13.8776 | -105.4579
13.8812 | -105.4508
13.8848 | -105.4403
13.8884 | -105.4304
13.8921 | -105.422
13.8957 | -105.4126
13.8993 | -105.404
13.9029 | -105.3951
13.9066 | -105.3895
13.9102 | -105.3821
13.9138 | -105.3715
13.9175 | -105.3615
13.9211 | -105.3535
13.9247 | -105.3484
13.9283 | -105.3464
13.932 | -105.3456
13.9356 | -105.3458
13.9392 | -105.3489
13.9429 | -105.3568
13.9465 | -105.3625
13.9501 | -105.3665
13.9537 | -105.3742
13.9574 | -105.3843
13.961 | -105.3874
13.9646 | -105.3846
13.9683 | -105.3859
13.9719 | -105.3898
13.9755 | -105.3902
13.9791 | -105.3884
13.9828 | -105.3865
13.9864 | -105.38
13.99 | -105.3741
13.9937 | -105.3668
13.9973 | -105.3555
14.0009 | -105.3425
14.0045 | -105.3262
14.0082 | -105.3081
14.0118 | -105.2897
14.0154 | -105.2676
14.019 | -105.2471
14.0227 | -105.225
14.0263 | -105.2003
14.0299 | -105.1722
14.0336 | -105.1472
14.0372 | -105.1227
14.0408 | -105.0988
14.0444 | -105.0723
14.0481 | -105.0461
14.0517 | -105.0269
14.0553 | -105.0148
14.059 | -105.0018
14.0626 | -104.9879
14.0662 | -104.9751
14.0698 | -104.9663
14.0735 | -104.9619
14.0771 | -104.9594
14.0807 | -104.9606
14.0844 | -104.9675
14.088 | -104.9736
14.0916 | -104.98
14.0952 | -104.9875
14.0989 | -104.9941
14.1025 | -105.0047
14.1061 | -105.0137
14.1098 | -105.0219
14.1134 | -105.0315
14.117 | -105.0431
14.1206 | -105.0522
14.1243 | -105.0596
14.1279 | -105.064
14.1315 | -105.0663
14.1351 | -105.0726
14.1388 | -105.0768
14.1424 | -105.0806
14.146 | -105.0828
14.1497 | -105.088
14.1533 | -105.0952
14.1569 | -105.0992
14.1605 | -105.1024
14.1642 | -105.1097
14.1678 | -105.1182
14.1714 | -105.1207
14.1751 | -105.1225
14.1787 | -105.1271
14.1823 | -105.1298
14.1859 | -105.1338
14.1896 | -105.1388
14.1932 | -105.1445
14.1968 | -105.1508
14.2005 | -105.1596
14.2041 | -105.1705
14.2077 | -105.1821
14.2113 | -105.1949
14.215 | -105.2077
14.2186 | -105.2224
14.2222 | -105.2378
14.2258 | -105.2537
14.2295 | -105.2718
14.2331 | -105.2903
14.2367 | -105.3065
14.2404 | -105.3235
14.244 | -105.3406
14.2476 | -105.3603
14.2512 | -105.3798
14.2549 | -105.3926
14.2585 | -105.4029
14.2621 | -105.4154
14.2658 | -105.4244
14.2694 | -105.4331
14.273 | -105.4399
14.2766 | -105.4443
14.2803 | -105.4462
14.2839 | -105.4455
14.2875 | -105.4437
14.2912 | -105.4406
14.2948 | -105.4351
14.2984 | -105.43
14.302 | -105.4258
14.3057 | -105.4161
14.3093 | -105.4051
14.3129 | -105.3976
14.3166 | -105.3888
14.3202 | -105.3768
14.3238 | -105.3667
14.3274 | -105.3566
14.3311 | -105.3509
14.3347 | -105.3444
14.3383 | -105.3372
14.342 | -105.331
14.3456 | -105.3263
14.3492 | -105.3257
14.3528 | -105.3265
14.3565 | -105.3285
14.3601 | -105.3313
14.3637 | -105.338
14.3673 | -105.347
14.371 | -105.3583
14.3746 | -105.3697
14.3782 | -105.3794
14.3819 | -105.3882
14.3855 | -105.3976
14.3891 | -105.4125
14.3927 | -105.4257
14.3964 | -105.4422
14.4 | -105.4574
14.4036 | -105.4723
14.4073 | -105.4869
14.4109 | -105.4991
14.4145 | -105.5118
14.4181 | -105.5212
14.4218 | -105.5316
14.4254 | -105.5422
14.429 | -105.5492
14.4327 | -105.5548
14.4363 | -105.5595
14.4399 | -105.5619
14.4435 | -105.5622
14.4472 | -105.5596
14.4508 | -105.5567
14.4544 | -105.5523
14.4581 | -105.546
14.4617 | -105.5392
14.4653 | -105.5338
14.4689 | -105.5256
14.4726 | -105.5175
14.4762 | -105.5089
14.4798 | -105.4973
14.4834 | -105.4868
14.4871 | -105.4777
14.4907 | -105.4686
14.4943 | -105.459
14.498 | -105.4502
14.5016 | -105.4432
14.5052 | -105.4354
14.5088 | -105.4282
14.5125 | -105.4231
14.5161 | -105.4206
14.5197 | -105.4195
14.5234 | -105.4181
14.527 | -105.4189
14.5306 | -105.4205
14.5342 | -105.4249
14.5379 | -105.4272
14.5415 | -105.4305
14.5451 | -105.4355
14.5488 | -105.4445
14.5524 | -105.455
14.556 | -105.4638
14.5596 | -105.4724
14.5633 | -105.4772
14.5669 | -105.4845
14.5705 | -105.4941
14.5741 | -105.5033
14.5778 | -105.5121
14.5814 | -105.5231
14.585 | -105.5356
14.5887 | -105.5461
14.5923 | -105.5531
14.5959 | -105.5536
14.5995 | -105.5551
14.6032 | -105.5537
14.6068 | -105.5507
14.6104 | -105.5473
14.6141 | -105.5434
14.6177 | -105.537
14.6213 | -105.5305
14.6249 | -105.5253
14.6286 | -105.5137
14.6322 | -105.5014
14.6358 | -105.4882
14.6395 | -105.4723
14.6431 | -105.4521
14.6467 | -105.4358
14.6503 | -105.4215
14.654 | -105.4063
14.6576 | -105.3929
14.6612 | -105.3762
14.6649 | -105.3612
14.6685 | -105.3488
14.6721 | -105.337
14.6757 | -105.3261
14.6794 | -105.3183
14.683 | -105.3116
14.6866 | -105.3019
14.6902 | -105.295
14.6939 | -105.2899
14.6975 | -105.2841
14.7011 | -105.2813
14.7048 | -105.2796
14.7084 | -105.2787
14.712 | -105.2798
14.7156 | -105.2818
14.7193 | -105.2832
14.7229 | -105.2836
14.7265 | -105.2845
14.7302 | -105.289
14.7338 | -105.2954
14.7374 | -105.2993
14.741 | -105.3069
14.7447 | -105.3137
14.7483 | -105.3243
14.7519 | -105.3333
14.7556 | -105.3398
14.7592 | -105.3485
14.7628 | -105.358
14.7664 | -105.3674
14.7701 | -105.3764
14.7737 | -105.3861
14.7773 | -105.396
14.781 | -105.4004
14.7846 | -105.408
14.7882 | -105.4148
14.7918 | -105.4189
14.7955 | -105.4226
14.7991 | -105.4264
14.8027 | -105.432
14.8063 | -105.4388
14.81 | -105.4426
14.8136 | -105.4458
14.8172 | -105.4489
14.8209 | -105.4523
14.8245 | -105.4558
14.8281 | -105.4589
14.8317 | -105.4612
14.8354 | -105.4658
14.839 | -105.471
14.8426 | -105.4707
14.8463 | -105.4716
14.8499 | -105.4755
14.8535 | -105.4812
14.8571 | -105.4822
14.8608 | -105.4822
14.8644 | -105.4833
14.868 | -105.4844
14.8717 | -105.4851
14.8753 | -105.4866
14.8789 | -105.4868
14.8825 | -105.4866
14.8862 | -105.4868
14.8898 | -105.4857
14.8934 | -105.4855
14.8971 | -105.4866
14.9007 | -105.4883
14.9043 | -105.4902
14.9079 | -105.4913
14.9116 | -105.4942
14.9152 | -105.4958
14.9188 | -105.4953
14.9224 | -105.496
14.9261 | -105.4957
14.9297 | -105.4964
14.9333 | -105.4992
14.937 | -105.5017
14.9406 | -105.5046
14.9442 | -105.5015
14.9478 | -105.4987
14.9515 | -105.4996
14.9551 | -105.497
14.9587 | -105.4938
14.9624 | -105.4863
14.966 | -105.4804
14.9696 | -105.4778
14.9732 | -105.4687
14.9769 | -105.4594
14.9805 | -105.448
14.9841 | -105.4373
14.9878 | -105.4283
14.9914 | -105.4167
14.995 | -105.4087
14.9986 | -105.3944
15.0023 | -105.3785
15.0059 | -105.3649
15.0095 | -105.3495
15.0132 | -105.3326
15.0168 | -105.3143
15.0204 | -105.2978
15.024 | -105.2825
15.0277 | -105.266
15.0313 | -105.2512
15.0349 | -105.2352
15.0385 | -105.2194
15.0422 | -105.2062
15.0458 | -105.1936
15.0494 | -105.1784
15.0531 | -105.1682
15.0567 | -105.1616
15.0603 | -105.1556
15.0639 | -105.1482
15.0676 | -105.1403
15.0712 | -105.1363
15.0748 | -105.1356
15.0785 | -105.1355
15.0821 | -105.134
15.0857 | -105.1342
15.0893 | -105.1398
15.093 | -105.1464
15.0966 | -105.1541
15.1002 | -105.1606
15.1039 | -105.1673
15.1075 | -105.174
15.1111 | -105.1828
15.1147 | -105.1954
15.1184 | -105.207
15.122 | -105.2176
15.1256 | -105.2303
15.1293 | -105.2421
15.1329 | -105.2569
15.1365 | -105.2704
15.1401 | -105.2818
15.1438 | -105.294
15.1474 | -105.3041
15.151 | -105.3139
15.1546 | -105.3224
15.1583 | -105.3332
15.1619 | -105.3424
15.1655 | -105.3485
15.1692 | -105.3555
15.1728 | -105.3633
15.1764 | -105.3675
15.18 | -105.3785
15.1837 | -105.3866
15.1873 | -105.3886
15.1909 | -105.3928
15.1946 | -105.3963
15.1982 | -105.3986
15.2018 | -105.4014
15.2054 | -105.4027
15.2091 | -105.4023
15.2127 | -105.4022
15.2163 | -105.4046
15.22 | -105.4064
15.2236 | -105.407
15.2272 | -105.4091
15.2308 | -105.4102
15.2345 | -105.4123
15.2381 | -105.414
15.2417 | -105.4118
15.2454 | -105.4088
15.249 | -105.4085
15.2526 | -105.412
15.2562 | -105.415
15.2599 | -105.4167
15.2635 | -105.4192
15.2671 | -105.423
15.2707 | -105.4247
15.2744 | -105.4251
15.278 | -105.4266
15.2816 | -105.4301
15.2853 | -105.4377
15.2889 | -105.4432
15.2925 | -105.4418
15.2961 | -105.4427
15.2998 | -105.4497
15.3034 | -105.453
15.307 | -105.4548
15.3107 | -105.4588
15.3143 | -105.4613
15.3179 | -105.465
15.3215 | -105.4695
15.3252 | -105.4717
15.3288 | -105.4713
15.3324 | -105.4717
15.3361 | -105.4719
15.3397 | -105.472
15.3433 | -105.4688
15.3469 | -105.4663
15.3506 | -105.4618
15.3542 | -105.4573
15.3578 | -105.452
15.3615 | -105.4446
15.3651 | -105.4379
15.3687 | -105.4308
15.3723 | -105.4226
15.376 | -105.411
15.3796 | -105.3988
15.3832 | -105.3882
15.3868 | -105.3763
15.3905 | -105.3616
15.3941 | -105.3468
15.3977 | -105.3357
15.4014 | -105.3236
15.405 | -105.3124
15.4086 | -105.3051
15.4122 | -105.2959
15.4159 | -105.288
15.4195 | -105.2826
15.4231 | -105.2776
15.4268 | -105.272
15.4304 | -105.2677
15.434 | -105.264
15.4376 | -105.2605
15.4413 | -105.2593
15.4449 | -105.2617
15.4485 | -105.2652
15.4522 | -105.2717
15.4558 | -105.279
15.4594 | -105.2851
15.463 | -105.2959
15.4667 | -105.3078
15.4703 | -105.3164
15.4739 | -105.3218
15.4776 | -105.3312
15.4812 | -105.3422
15.4848 | -105.3523
15.4884 | -105.3626
15.4921 | -105.3722
15.4957 | -105.3843
15.4993 | -105.3931
15.5029 | -105.4028
15.5066 | -105.4086
15.5102 | -105.411
15.5138 | -105.4195
15.5175 | -105.4253
15.5211 | -105.4275
15.5247 | -105.4314
15.5283 | -105.4314
15.532 | -105.4291
15.5356 | -105.4271
15.5392 | -105.427
15.5429 | -105.4261
15.5465 | -105.4252
15.5501 | -105.4231
15.5537 | -105.4145
15.5574 | -105.4
15.561 | -105.3855
15.5646 | -105.3718
15.5683 | -105.3607
15.5719 | -105.3484
15.5755 | -105.335
15.5791 | -105.3186
15.5828 | -105.3024
15.5864 | -105.2846
15.59 | -105.2679
15.5937 | -105.2503
15.5973 | -105.2331
15.6009 | -105.2188
15.6045 | -105.2026
15.6082 | -105.184
15.6118 | -105.1669
15.6154 | -105.1497
15.619 | -105.1332
15.6227 | -105.1166
15.6263 | -105.1003
15.6299 | -105.0858
15.6336 | -105.0752
15.6372 | -105.066
15.6408 | -105.0569
15.6444 | -105.049
15.6481 | -105.0438
15.6517 | -105.0401
15.6553 | -105.0383
15.659 | -105.0405
15.6626 | -105.0419
15.6662 | -105.0382
15.6698 | -105.0376
15.6735 | -105.0415
15.6771 | -105.0476
15.6807 | -105.0552
15.6844 | -105.064
15.688 | -105.0727
15.6916 | -105.0816
15.6952 | -105.0955
15.6989 | -105.1068
15.7025 | -105.1211
15.7061 | -105.1335
15.7098 | -105.1458
15.7134 | -105.1578
15.717 | -105.1693
15.7206 | -105.1801
15.7243 | -105.191
15.7279 | -105.201
15.7315 | -105.2088
15.7351 | -105.2163
15.7388 | -105.2259
15.7424 | -105.2325
15.746 | -105.2382
15.7497 | -105.2443
15.7533 | -105.2513
15.7569 | -105.2562
15.7605 | -105.2603
15.7642 | -105.2636
15.7678 | -105.2657
15.7714 | -105.2672
15.7751 | -105.2687
15.7787 | -105.2717
15.7823 | -105.2785
15.7859 | -105.283
15.7896 | -105.2839
15.7932 | -105.2799
15.7968 | -105.2756
15.8005 | -105.2742
15.8041 | -105.2754
15.8077 | -105.2749
15.8113 | -105.2765
15.815 | -105.2748
15.8186 | -105.2707
15.8222 | -105.2659
15.8259 | -105.2597
15.8295 | -105.254
15.8331 | -105.2463
15.8367 | -105.241
15.8404 | -105.2375
15.844 | -105.2318
15.8476 | -105.2255
15.8512 | -105.2188
15.8549 | -105.2079
15.8585 | -105.1967
15.8621 | -105.1839
15.8658 | -105.1725
15.8694 | -105.1612
15.873 | -105.1494
15.8766 | -105.1367
15.8803 | -105.1217
15.8839 | -105.1075
15.8875 | -105.0942
15.8912 | -105.0826
15.8948 | -105.0677
15.8984 | -105.0488
15.902 | -105.0304
15.9057 | -105.013
15.9093 | -104.9964
15.9129 | -104.9809
15.9166 | -104.9648
15.9202 | -104.9464
15.9238 | -104.9287
15.9274 | -104.9111
15.9311 | -104.8907
15.9347 | -104.8724
15.9383 | -104.8559
15.942 | -104.8419
15.9456 | -104.8295
15.9492 | -104.8154
15.9528 | -104.8028
15.9565 | -104.7897
15.9601 | -104.7799
15.9637 | -104.7713
15.9673 | -104.7592
15.971 | -104.7502
15.9746 | -104.7383
15.9782 | -104.7268
15.9819 | -104.7172
15.9855 | -104.7095
15.9891 | -104.7031
15.9927 | -104.6958
15.9964 | -104.689
16 | -104.6838
16.0036 | -104.6788
16.0073 | -104.6732
16.0109 | -104.6686
16.0145 | -104.666
16.0181 | -104.6628
16.0218 | -104.6599
16.0254 | -104.6582
16.029 | -104.6551
16.0327 | -104.6506
16.0363 | -104.6471
16.0399 | -104.6424
16.0435 | -104.6385
16.0472 | -104.6385
16.0508 | -104.6397
16.0544 | -104.6379
16.0581 | -104.6341
16.0617 | -104.634
16.0653 | -104.6297
16.0689 | -104.623
16.0726 | -104.6176
16.0762 | -104.6128
16.0798 | -104.6109
16.0834 | -104.6092
16.0871 | -104.6055
16.0907 | -104.6
16.0943 | -104.5934
16.098 | -104.5877
16.1016 | -104.5834
16.1052 | -104.5791
16.1088 | -104.5752
16.1125 | -104.5706
16.1161 | -104.5682
16.1197 | -104.567
16.1234 | -104.5672
16.127 | -104.5678
16.1306 | -104.5693
16.1342 | -104.5713
16.1379 | -104.574
16.1415 | -104.5795
16.1451 | -104.5855
16.1488 | -104.5913
16.1524 | -104.5985
16.156 | -104.6084
16.1596 | -104.62
16.1633 | -104.6299
16.1669 | -104.6411
16.1705 | -104.6491
16.1742 | -104.6557
16.1778 | -104.6658
16.1814 | -104.6765
16.185 | -104.6864
16.1887 | -104.6949
16.1923 | -104.7057
16.1959 | -104.7141
16.1995 | -104.7202
16.2032 | -104.7321
16.2068 | -104.7418
16.2104 | -104.7496
16.2141 | -104.7557
16.2177 | -104.762
16.2213 | -104.7704
16.2249 | -104.7778
16.2286 | -104.7853
16.2322 | -104.7883
16.2358 | -104.7955
16.2395 | -104.8043
16.2431 | -104.8116
16.2467 | -104.8179
16.2503 | -104.8247
16.254 | -104.8322
16.2576 | -104.8379
16.2612 | -104.8452
16.2649 | -104.8525
16.2685 | -104.8566
16.2721 | -104.8613
16.2757 | -104.8677
16.2794 | -104.8731
16.283 | -104.8759
16.2866 | -104.879
16.2902 | -104.8835
16.2939 | -104.8858
16.2975 | -104.8888
16.3011 | -104.8917
16.3048 | -104.8934
16.3084 | -104.8955
16.312 | -104.8971
16.3156 | -104.9008
16.3193 | -104.9052
16.3229 | -104.9115
16.3265 | -104.9168
16.3302 | -104.9215
16.3338 | -104.9268
16.3374 | -104.9336
16.341 | -104.94
16.3447 | -104.9474
16.3483 | -104.9541
16.3519 | -104.9619
16.3556 | -104.9722
16.3592 | -104.983
16.3628 | -104.9927
16.3664 | -105.004
16.3701 | -105.0132
16.3737 | -105.0208
16.3773 | -105.0285
16.381 | -105.0349
16.3846 | -105.0433
16.3882 | -105.0507
16.3918 | -105.058
16.3955 | -105.0661
16.3991 | -105.071
16.4027 | -105.077
16.4063 | -105.0833
16.41 | -105.0869
16.4136 | -105.0892
16.4172 | -105.0926
16.4209 | -105.097
16.4245 | -105.0981
16.4281 | -105.1024
16.4317 | -105.1083
16.4354 | -105.1091
16.439 | -105.1105
16.4426 | -105.1137
16.4463 | -105.1186
16.4499 | -105.1245
16.4535 | -105.1275
16.4571 | -105.1299
16.4608 | -105.1341
16.4644 | -105.1383
16.468 | -105.1459
16.4717 | -105.1515
16.4753 | -105.1566
16.4789 | -105.1609
16.4825 | -105.1641
16.4862 | -105.1692
16.4898 | -105.1746
16.4934 | -105.181
16.4971 | -105.1871
16.5007 | -105.194
16.5043 | -105.2026
16.5079 | -105.2117
16.5116 | -105.2185
16.5152 | -105.2242
16.5188 | -105.2331
16.5224 | -105.2417
16.5261 | -105.2511
16.5297 | -105.2594
16.5333 | -105.2673
16.537 | -105.2762
16.5406 | -105.2879
16.5442 | -105.2976
16.5478 | -105.3037
16.5515 | -105.31
16.5551 | -105.3171
16.5587 | -105.326
16.5624 | -105.3355
16.566 | -105.3436
16.5696 | -105.3499
16.5732 | -105.3551
16.5769 | -105.3619
16.5805 | -105.3699
16.5841 | -105.3775
16.5878 | -105.3833
16.5914 | -105.387
16.595 | -105.3917
16.5986 | -105.396
16.6023 | -105.3998
16.6059 | -105.404
16.6095 | -105.4104
16.6132 | -105.4187
16.6168 | -105.4276
16.6204 | -105.4336
16.624 | -105.4394
16.6277 | -105.4448
16.6313 | -105.4523
16.6349 | -105.4622
16.6385 | -105.4698
16.6422 | -105.4799
16.6458 | -105.491
16.6494 | -105.5041
16.6531 | -105.5165
16.6567 | -105.5275
16.6603 | -105.5381
16.6639 | -105.5512
16.6676 | -105.5636
16.6712 | -105.5785
16.6748 | -105.5924
16.6785 | -105.6044
16.6821 | -105.6174
16.6857 | -105.628
16.6893 | -105.6394
16.693 | -105.6512
16.6966 | -105.6617
16.7002 | -105.6733
16.7039 | -105.6843
16.7075 | -105.6925
16.7111 | -105.7009
16.7147 | -105.708
16.7184 | -105.7156
16.722 | -105.7187
16.7256 | -105.7181
16.7293 | -105.7134
16.7329 | -105.7107
16.7365 | -105.7058
16.7401 | -105.7014
16.7438 | -105.6961
16.7474 | -105.6791
16.751 | -105.6701
16.7546 | -105.6617
16.7583 | -105.6492
16.7619 | -105.6381
16.7655 | -105.6242
16.7692 | -105.6108
16.7728 | -105.5976
16.7764 | -105.5806
16.78 | -105.5615
16.7837 | -105.5438
16.7873 | -105.5274
16.7909 | -105.5095
16.7946 | -105.4925
16.7982 | -105.4744
16.8018 | -105.4572
16.8054 | -105.44
16.8091 | -105.4245
16.8127 | -105.4101
16.8163 | -105.3934
16.82 | -105.3761
16.8236 | -105.3599
16.8272 | -105.3468
16.8308 | -105.3349
16.8345 | -105.3259
16.8381 | -105.3191
16.8417 | -105.3129
16.8454 | -105.3088
16.849 | -105.3046
16.8526 | -105.301
16.8562 | -105.3001
16.8599 | -105.3015
16.8635 | -105.3041
16.8671 | -105.3067
16.8707 | -105.3096
16.8744 | -105.3137
16.878 | -105.3212
16.8816 | -105.3268
16.8853 | -105.3341
16.8889 | -105.3416
16.8925 | -105.3499
16.8961 | -105.3588
16.8998 | -105.37
16.9034 | -105.3785
16.907 | -105.3843
16.9107 | -105.392
16.9143 | -105.4012
16.9179 | -105.4108
16.9215 | -105.4184
16.9252 | -105.4256
16.9288 | -105.435
16.9324 | -105.4426
16.9361 | -105.449
16.9397 | -105.4547
16.9433 | -105.4574
16.9469 | -105.4587
16.9506 | -105.4593
16.9542 | -105.4614
16.9578 | -105.4641
16.9615 | -105.4662
16.9651 | -105.4665
16.9687 | -105.4706
16.9723 | -105.4683
16.976 | -105.468
16.9796 | -105.4647
16.9832 | -105.4626
16.9868 | -105.463
16.9905 | -105.463
16.9941 | -105.4626
16.9977 | -105.4607
17.0014 | -105.4568
17.005 | -105.4549
17.0086 | -105.4524
17.0122 | -105.449
17.0159 | -105.4459
17.0195 | -105.4459
17.0231 | -105.4459
17.0268 | -105.4448
17.0304 | -105.4445
17.034 | -105.445
17.0376 | -105.4437
17.0413 | -105.441
17.0449 | -105.4398
17.0485 | -105.4389
17.0522 | -105.4369
17.0558 | -105.4367
17.0594 | -105.4354
17.063 | -105.4317
17.0667 | -105.4303
17.0703 | -105.4299
17.0739 | -105.4293
17.0776 | -105.4267
17.0812 | -105.4242
17.0848 | -105.4225
17.0884 | -105.4206
17.0921 | -105.4192
17.0957 | -105.4172
17.0993 | -105.4163
17.1029 | -105.4143
17.1066 | -105.4124
17.1102 | -105.4134
17.1138 | -105.4124
17.1175 | -105.4116
17.1211 | -105.4127
17.1247 | -105.4151
17.1283 | -105.4165
17.132 | -105.4158
17.1356 | -105.4175
17.1392 | -105.4178
17.1429 | -105.4141
17.1465 | -105.4118
17.1501 | -105.412
17.1537 | -105.411
17.1574 | -105.409
17.161 | -105.4071
17.1646 | -105.4047
17.1683 | -105.4028
17.1719 | -105.3993
17.1755 | -105.3982
17.1791 | -105.3957
17.1828 | -105.3929
17.1864 | -105.3917
17.19 | -105.3916
17.1937 | -105.3903
17.1973 | -105.3851
17.2009 | -105.3769
17.2045 | -105.3714
17.2082 | -105.366
17.2118 | -105.3543
17.2154 | -105.3447
17.219 | -105.3354
17.2227 | -105.3259
17.2263 | -105.3134
17.2299 | -105.3005
17.2336 | -105.2899
17.2372 | -105.2779
17.2408 | -105.2646
17.2444 | -105.2552
17.2481 | -105.2438
17.2517 | -105.2323
17.2553 | -105.2219
17.259 | -105.2097
17.2626 | -105.1949
17.2662 | -105.1801
17.2698 | -105.1706
17.2735 | -105.1603
17.2771 | -105.1511
17.2807 | -105.1434
17.2844 | -105.1346
17.288 | -105.129
17.2916 | -105.1232
17.2952 | -105.1179
17.2989 | -105.1152
17.3025 | -105.1107
17.3061 | -105.1124
17.3098 | -105.1137
17.3134 | -105.1128
17.317 | -105.1116
17.3206 | -105.113
17.3243 | -105.1167
17.3279 | -105.1194
17.3315 | -105.1234
17.3351 | -105.1281
17.3388 | -105.1332
17.3424 | -105.14
17.346 | -105.1456
17.3497 | -105.1518
17.3533 | -105.1582
17.3569 | -105.166
17.3605 | -105.1747
17.3642 | -105.1827
17.3678 | -105.189
17.3714 | -105.1946
17.3751 | -105.199
17.3787 | -105.2049
17.3823 | -105.2102
17.3859 | -105.2167
17.3896 | -105.2253
17.3932 | -105.2305
17.3968 | -105.2369
17.4005 | -105.2413
17.4041 | -105.2441
17.4077 | -105.2455
17.4113 | -105.2461
17.415 | -105.2453
17.4186 | -105.2428
17.4222 | -105.24
17.4259 | -105.2356
17.4295 | -105.2276
17.4331 | -105.218
17.4367 | -105.2091
17.4404 | -105.1994
17.444 | -105.1901
17.4476 | -105.1795
17.4512 | -105.1647
17.4549 | -105.1496
17.4585 | -105.1343
17.4621 | -105.118
17.4658 | -105.0991
17.4694 | -105.0811
17.473 | -105.0631
17.4766 | -105.0417
17.4803 | -105.0221
17.4839 | -105.0016
17.4875 | -104.9814
17.4912 | -104.9635
17.4948 | -104.9451
17.4984 | -104.9276
17.502 | -104.9106
17.5057 | -104.8951
17.5093 | -104.8809
17.5129 | -104.8685
17.5166 | -104.8568
17.5202 | -104.8466
17.5238 | -104.8368
17.5274 | -104.8274
17.5311 | -104.8191
17.5347 | -104.8144
17.5383 | -104.8107
17.5419 | -104.8084
17.5456 | -104.8078
17.5492 | -104.8071
17.5528 | -104.8118
17.5565 | -104.8158
17.5601 | -104.8227
17.5637 | -104.8306
17.5673 | -104.8405
17.571 | -104.8511
17.5746 | -104.8617
17.5782 | -104.874
17.5819 | -104.8876
17.5855 | -104.904
17.5891 | -104.9203
17.5927 | -104.9364
17.5964 | -104.9537
17.6 | -104.9726
17.6036 | -104.9914
17.6073 | -105.0088
17.6109 | -105.0277
17.6145 | -105.0466
17.6181 | -105.0653
17.6218 | -105.0859
17.6254 | -105.1077
17.629 | -105.1273
17.6327 | -105.1436
17.6363 | -105.1586
17.6399 | -105.173
17.6435 | -105.1883
17.6472 | -105.2016
17.6508 | -105.2131
17.6544 | -105.2234
17.658 | -105.231
17.6617 | -105.2409
17.6653 | -105.2489
17.6689 | -105.2555
17.6726 | -105.2581
17.6762 | -105.2602
17.6798 | -105.2614
17.6834 | -105.2634
17.6871 | -105.2632
17.6907 | -105.2613
17.6943 | -105.2567
17.698 | -105.2476
17.7016 | -105.2399
17.7052 | -105.2333
17.7088 | -105.2261
17.7125 | -105.2164
17.7161 | -105.2075
17.7197 | -105.1963
17.7234 | -105.1838
17.727 | -105.1738
17.7306 | -105.1617
17.7342 | -105.1512
17.7379 | -105.1408
17.7415 | -105.128
17.7451 | -105.1138
17.7488 | -105.1006
17.7524 | -105.0891
17.756 | -105.0758
17.7596 | -105.0602
17.7633 | -105.0455
17.7669 | -105.0307
17.7705 | -105.0158
17.7741 | -105.0013
17.7778 | -104.9894
17.7814 | -104.9776
17.785 | -104.9654
17.7887 | -104.9549
17.7923 | -104.9451
17.7959 | -104.9388
17.7995 | -104.9307
17.8032 | -104.9222
17.8068 | -104.9111
17.8104 | -104.9024
17.8141 | -104.8944
17.8177 | -104.8889
17.8213 | -104.8829
17.8249 | -104.8794
17.8286 | -104.8796
17.8322 | -104.8789
17.8358 | -104.8789
17.8395 | -104.8793
17.8431 | -104.8786
17.8467 | -104.8799
17.8503 | -104.8814
17.854 | -104.885
17.8576 | -104.8873
17.8612 | -104.8902
17.8649 | -104.8948
17.8685 | -104.8976
17.8721 | -104.8993
17.8757 | -104.9006
17.8794 | -104.9023
17.883 | -104.9064
17.8866 | -104.9097
17.8902 | -104.9143
17.8939 | -104.917
17.8975 | -104.9176
17.9011 | -104.9162
17.9048 | -104.9168
17.9084 | -104.9149
17.912 | -104.9103
17.9156 | -104.9042
17.9193 | -104.8991
17.9229 | -104.8945
17.9265 | -104.8894
17.9302 | -104.8842
17.9338 | -104.8768
17.9374 | -104.8661
17.941 | -104.8552
17.9447 | -104.8442
17.9483 | -104.8359
17.9519 | -104.8299
17.9556 | -104.8205
17.9592 | -104.8074
17.9628 | -104.7941
17.9664 | -104.7829
17.9701 | -104.7703
17.9737 | -104.7592
17.9773 | -104.7514
17.981 | -104.7434
17.9846 | -104.7362
17.9882 | -104.7296
17.9918 | -104.724
17.9955 | -104.7178
17.9991 | -104.7151
18.0027 | -104.7124
18.0063 | -104.7107
18.01 | -104.7133
18.0136 | -104.7197
18.0172 | -104.7261
18.0209 | -104.7293
18.0245 | -104.734
18.0281 | -104.7418
18.0317 | -104.7469
18.0354 | -104.7537
18.039 | -104.7631
18.0426 | -104.7727
18.0463 | -104.7845
18.0499 | -104.7958
18.0535 | -104.8072
18.0571 | -104.8191
18.0608 | -104.8322
18.0644 | -104.8446
18.068 | -104.8566
18.0717 | -104.8689
18.0753 | -104.8808
18.0789 | -104.8906
18.0825 | -104.8987
18.0862 | -104.9063
18.0898 | -104.9149
18.0934 | -104.9202
18.0971 | -104.9252
18.1007 | -104.9296
18.1043 | -104.9321
18.1079 | -104.9355
18.1116 | -104.9367
18.1152 | -104.9388
18.1188 | -104.9416
18.1224 | -104.9411
18.1261 | -104.9407
18.1297 | -104.9406
18.1333 | -104.9372
18.137 | -104.9317
18.1406 | -104.9249
18.1442 | -104.9221
18.1478 | -104.9174
18.1515 | -104.9126
18.1551 | -104.907
18.1587 | -104.901
18.1624 | -104.8941
18.166 | -104.8899
18.1696 | -104.8867
18.1732 | -104.8802
18.1769 | -104.8706
18.1805 | -104.8603
18.1841 | -104.8497
18.1878 | -104.8404
18.1914 | -104.8346
18.195 | -104.8279
18.1986 | -104.8208
18.2023 | -104.8154
18.2059 | -104.8082
18.2095 | -104.8011
18.2132 | -104.7938
18.2168 | -104.7888
18.2204 | -104.7841
18.224 | -104.7766
18.2277 | -104.7693
18.2313 | -104.7627
18.2349 | -104.755
18.2385 | -104.7453
18.2422 | -104.7358
18.2458 | -104.7279
18.2494 | -104.7212
18.2531 | -104.7172
18.2567 | -104.7112
18.2603 | -104.7049
18.2639 | -104.6961
18.2676 | -104.6871
18.2712 | -104.676
18.2748 | -104.6641
18.2785 | -104.6549
18.2821 | -104.6429
18.2857 | -104.6322
18.2893 | -104.6243
18.293 | -104.6139
18.2966 | -104.6022
18.3002 | -104.5907
18.3039 | -104.5792
18.3075 | -104.5685
18.3111 | -104.5591
18.3147 | -104.5465
18.3184 | -104.5321
18.322 | -104.5214
18.3256 | -104.5114
18.3293 | -104.4989
18.3329 | -104.4858
18.3365 | -104.4738
18.3401 | -104.4635
18.3438 | -104.4527
18.3474 | -104.4405
18.351 | -104.4293
18.3546 | -104.4187
18.3583 | -104.4083
18.3619 | -104.3988
18.3655 | -104.3897
18.3692 | -104.3812
18.3728 | -104.3711
18.3764 | -104.3619
18.38 | -104.3487
18.3837 | -104.3398
18.3873 | -104.3324
18.3909 | -104.3255
18.3946 | -104.3189
18.3982 | -104.3137
18.4018 | -104.3071
18.4054 | -104.3003
18.4091 | -104.2933
18.4127 | -104.2867
18.4163 | -104.2825
18.42 | -104.278
18.4236 | -104.2741
18.4272 | -104.2713
18.4308 | -104.2669
18.4345 | -104.2625
18.4381 | -104.2582
18.4417 | -104.2539
18.4454 | -104.2514
18.449 | -104.2463
18.4526 | -104.2389
18.4562 | -104.2336
18.4599 | -104.2278
18.4635 | -104.22
18.4671 | -104.2138
18.4707 | -104.2072
18.4744 | -104.2014
18.478 | -104.1932
18.4816 | -104.1851
18.4853 | -104.1787
18.4889 | -104.17
18.4925 | -104.1612
18.4961 | -104.1497
18.4998 | -104.1412
18.5034 | -104.1327
18.507 | -104.1234
18.5107 | -104.1145
18.5143 | -104.1022
18.5179 | -104.0915
18.5215 | -104.0837
18.5252 | -104.0751
18.5288 | -104.066
18.5324 | -104.0568
18.5361 | -104.0475
18.5397 | -104.0398
18.5433 | -104.0313
18.5469 | -104.0202
18.5506 | -104.0077
18.5542 | -103.9951
18.5578 | -103.9831
18.5615 | -103.9692
18.5651 | -103.9516
18.5687 | -103.9349
18.5723 | -103.9187
18.576 | -103.9042
18.5796 | -103.8903
18.5832 | -103.8714
18.5868 | -103.8516
18.5905 | -103.832
18.5941 | -103.8127
18.5977 | -103.7937
18.6014 | -103.7735
18.605 | -103.7526
18.6086 | -103.7317
18.6122 | -103.7123
18.6159 | -103.6924
18.6195 | -103.6739
18.6231 | -103.6529
18.6268 | -103.6311
18.6304 | -103.611
18.634 | -103.5884
18.6376 | -103.5668
18.6413 | -103.5456
18.6449 | -103.5282
18.6485 | -103.5081
18.6522 | -103.4885
18.6558 | -103.4695
18.6594 | -103.4493
18.663 | -103.4304
18.6667 | -103.4131
18.6703 | -103.3962
18.6739 | -103.3807
18.6776 | -103.3662
18.6812 | -103.3519
18.6848 | -103.3364
18.6884 | -103.3213
18.6921 | -103.3052
18.6957 | -103.2942
18.6993 | -103.2826
18.7029 | -103.2715
18.7066 | -103.2601
18.7102 | -103.2473
18.7138 | -103.2371
18.7175 | -103.2288
18.7211 | -103.219
18.7247 | -103.2093
18.7283 | -103.2007
18.732 | -103.1918
18.7356 | -103.1833
18.7392 | -103.1758
18.7429 | -103.1694
18.7465 | -103.1648
18.7501 | -103.16
18.7537 | -103.1554
18.7574 | -103.1495
18.761 | -103.1466
18.7646 | -103.141
18.7683 | -103.1342
18.7719 | -103.1281
18.7755 | -103.123
18.7791 | -103.12
18.7828 | -103.1163
18.7864 | -103.1099
18.79 | -103.1049
18.7937 | -103.1006
18.7973 | -103.0967
18.8009 | -103.0924
18.8045 | -103.0853
18.8082 | -103.0781
18.8118 | -103.0701
18.8154 | -103.0642
18.819 | -103.0569
18.8227 | -103.0473
18.8263 | -103.0412
18.8299 | -103.0361
18.8336 | -103.0285
18.8372 | -103.019
18.8408 | -103.0092
18.8444 | -102.998
18.8481 | -102.9821
18.8517 | -102.9666
18.8553 | -102.9519
18.859 | -102.9374
18.8626 | -102.923
18.8662 | -102.9073
18.8698 | -102.8909
18.8735 | -102.8748
18.8771 | -102.8587
18.8807 | -102.8409
18.8844 | -102.8238
18.888 | -102.8053
18.8916 | -102.7903
18.8952 | -102.7726
18.8989 | -102.7534
18.9025 | -102.7345
18.9061 | -102.7156
18.9098 | -102.6965
18.9134 | -102.6772
18.917 | -102.6587
18.9206 | -102.6414
18.9243 | -102.6248
18.9279 | -102.6077
18.9315 | -102.5902
18.9351 | -102.574
18.9388 | -102.5557
18.9424 | -102.5371
18.946 | -102.5217
18.9497 | -102.5068
18.9533 | -102.4951
18.9569 | -102.4822
18.9605 | -102.4698
18.9642 | -102.4565
18.9678 | -102.4477
18.9714 | -102.4403
18.9751 | -102.4331
18.9787 | -102.4277
18.9823 | -102.4213
18.9859 | -102.4151
18.9896 | -102.4102
18.9932 | -102.4054
18.9968 | -102.4021
19.0005 | -102.4013
19.0041 | -102.4025
19.0077 | -102.4031
19.0113 | -102.403
19.015 | -102.4011
19.0186 | -102.3991
19.0222 | -102.3998
19.0258 | -102.4035
19.0295 | -102.4056
19.0331 | -102.4077
19.0367 | -102.4109
19.0404 | -102.414
19.044 | -102.4143
19.0476 | -102.4169
19.0512 | -102.4186
19.0549 | -102.4208
19.0585 | -102.4223
19.0621 | -102.4234
19.0658 | -102.4257
19.0694 | -102.4268
19.073 | -102.4286
19.0766 | -102.4304
19.0803 | -102.4319
19.0839 | -102.432
19.0875 | -102.4314
19.0912 | -102.4309
19.0948 | -102.4285
19.0984 | -102.4266
19.102 | -102.4255
19.1057 | -102.423
19.1093 | -102.4167
19.1129 | -102.4099
19.1166 | -102.4066
19.1202 | -102.4072
19.1238 | -102.4073
19.1274 | -102.4072
19.1311 | -102.4041
19.1347 | -102.3991
19.1383 | -102.3968
19.142 | -102.3929
19.1456 | -102.3896
19.1492 | -102.3875
19.1528 | -102.3857
19.1565 | -102.3846
19.1601 | -102.3845
19.1637 | -102.3808
19.1673 | -102.3814
19.171 | -102.3822
19.1746 | -102.3843
19.1782 | -102.3866
19.1819 | -102.3877
19.1855 | -102.3874
19.1891 | -102.3891
19.1927 | -102.3889
19.1964 | -102.3884
19.2 | -102.3891
19.2036 | -102.3868
19.2073 | -102.3843
19.2109 | -102.3843
19.2145 | -102.3862
19.2181 | -102.3835
19.2218 | -102.3793
19.2254 | -102.3744
19.229 | -102.3695
19.2327 | -102.3678
19.2363 | -102.3656
19.2399 | -102.3608
19.2435 | -102.355
19.2472 | -102.3498
19.2508 | -102.3447
19.2544 | -102.3412
19.2581 | -102.3354
19.2617 | -102.3297
19.2653 | -102.3271
19.2689 | -102.3246
19.2726 | -102.3199
19.2762 | -102.3142
19.2798 | -102.3117
19.2834 | -102.3106
19.2871 | -102.3124
19.2907 | -102.3155
19.2943 | -102.3181
19.298 | -102.3191
19.3016 | -102.3198
19.3052 | -102.3233
19.3088 | -102.3266
19.3125 | -102.3317
19.3161 | -102.34
19.3197 | -102.3457
19.3234 | -102.3525
19.327 | -102.3617
19.3306 | -102.3724
19.3342 | -102.3822
19.3379 | -102.3922
19.3415 | -102.4033
19.3451 | -102.4161
19.3488 | -102.4308
19.3524 | -102.4446
19.356 | -102.4579
19.3596 | -102.4738
19.3633 | -102.4893
19.3669 | -102.5062
19.3705 | -102.5217
19.3742 | -102.5344
19.3778 | -102.5495
19.3814 | -102.5658
19.385 | -102.5837
19.3887 | -102.6015
19.3923 | -102.6171
19.3959 | -102.6325
19.3995 | -102.6481
19.4032 | -102.6606
19.4068 | -102.6745
19.4104 | -102.6896
19.4141 | -102.7025
19.4177 | -102.7156
19.4213 | -102.7266
19.4249 | -102.736
19.4286 | -102.7466
19.4322 | -102.7571
19.4358 | -102.7658
19.4395 | -102.7748
19.4431 | -102.783
19.4467 | -102.7922
19.4503 | -102.8006
19.454 | -102.8072
19.4576 | -102.8142
19.4612 | -102.8206
19.4649 | -102.8236
19.4685 | -102.8273
19.4721 | -102.8307
19.4757 | -102.8365
19.4794 | -102.8416
19.483 | -102.8433
19.4866 | -102.84
19.4902 | -102.8381
19.4939 | -102.8375
19.4975 | -102.837
19.5011 | -102.8383
19.5048 | -102.8386
19.5084 | -102.8345
19.512 | -102.8286
19.5156 | -102.8221
19.5193 | -102.8185
19.5229 | -102.8166
19.5265 | -102.814
19.5302 | -102.8102
19.5338 | -102.81
19.5374 | -102.8101
19.541 | -102.8066
19.5447 | -102.8017
19.5483 | -102.8
19.5519 | -102.7989
19.5556 | -102.7958
19.5592 | -102.7937
19.5628 | -102.7907
19.5664 | -102.7892
19.5701 | -102.7876
19.5737 | -102.7863
19.5773 | -102.7875
19.581 | -102.7888
19.5846 | -102.7902
19.5882 | -102.7902
19.5918 | -102.7872
19.5955 | -102.7853
19.5991 | -102.7864
19.6027 | -102.7879
19.6063 | -102.7865
19.61 | -102.7891
19.6136 | -102.7906
19.6172 | -102.7926
19.6209 | -102.7951
19.6245 | -102.7968
19.6281 | -102.7996
19.6317 | -102.8037
19.6354 | -102.807
19.639 | -102.8108
19.6426 | -102.8176
19.6463 | -102.8233
19.6499 | -102.8267
19.6535 | -102.832
19.6571 | -102.8348
19.6608 | -102.8384
19.6644 | -102.8428
19.668 | -102.8454
19.6717 | -102.8498
19.6753 | -102.8545
19.6789 | -102.8589
19.6825 | -102.8653
19.6862 | -102.8712
19.6898 | -102.8766
19.6934 | -102.8823
19.6971 | -102.889
19.7007 | -102.8945
19.7043 | -102.9005
19.7079 | -102.9054
19.7116 | -102.9089
19.7152 | -102.9143
19.7188 | -102.919
19.7224 | -102.922
19.7261 | -102.9222
19.7297 | -102.9234
19.7333 | -102.9273
19.737 | -102.932
19.7406 | -102.9365
19.7442 | -102.9402
19.7478 | -102.9442
19.7515 | -102.9477
19.7551 | -102.9483
19.7587 | -102.9526
19.7624 | -102.9577
19.766 | -102.9666
19.7696 | -102.9719
19.7732 | -102.9761
19.7769 | -102.9819
19.7805 | -102.9864
19.7841 | -102.9911
19.7878 | -102.9966
19.7914 | -103.0039
19.795 | -103.0109
19.7986 | -103.017
19.8023 | -103.0212
19.8059 | -103.0265
19.8095 | -103.0328
19.8132 | -103.0361
19.8168 | -103.0412
19.8204 | -103.0452
19.824 | -103.0511
19.8277 | -103.0562
19.8313 | -103.0596
19.8349 | -103.0618
19.8385 | -103.0666
19.8422 | -103.07
19.8458 | -103.0727
19.8494 | -103.0768
19.8531 | -103.0801
19.8567 | -103.0816
19.8603 | -103.0837
19.8639 | -103.0869
19.8676 | -103.0902
19.8712 | -103.0955
19.8748 | -103.1024
19.8785 | -103.1058
19.8821 | -103.1069
19.8857 | -103.1098
19.8893 | -103.1125
19.893 | -103.1136
19.8966 | -103.1161
19.9002 | -103.1184
19.9039 | -103.1192
19.9075 | -103.1209
19.9111 | -103.1247
19.9147 | -103.1285
19.9184 | -103.1318
19.922 | -103.1364
19.9256 | -103.1391
19.9293 | -103.1415
19.9329 | -103.1493
19.9365 | -103.1576
19.9401 | -103.1627
19.9438 | -103.1687
19.9474 | -103.1754
19.951 | -103.1821
19.9546 | -103.1874
19.9583 | -103.193
19.9619 | -103.1985
19.9655 | -103.2035
19.9692 | -103.2099
19.9728 | -103.2156
19.9764 | -103.2204
19.98 | -103.2265
19.9837 | -103.2333
19.9873 | -103.2387
19.9909 | -103.2447
19.9946 | -103.2518
19.9982 | -103.2576
20.0018 | -103.2637
20.0054 | -103.2684
20.0091 | -103.2718
20.0127 | -103.2742
20.0163 | -103.2756
20.02 | -103.2787
20.0236 | -103.2807
20.0272 | -103.2814
20.0308 | -103.2825
20.0345 | -103.2826
20.0381 | -103.2843
20.0417 | -103.284
20.0454 | -103.283
20.049 | -103.28
20.0526 | -103.2759
20.0562 | -103.2732
20.0599 | -103.2699
20.0635 | -103.2652
20.0671 | -103.2634
20.0707 | -103.2602
20.0744 | -103.2539
20.078 | -103.2497
20.0816 | -103.2413
20.0853 | -103.2324
20.0889 | -103.225
20.0925 | -103.2168
20.0961 | -103.2094
20.0998 | -103.1981
20.1034 | -103.1881
20.107 | -103.1772
20.1107 | -103.1657
20.1143 | -103.1559
20.1179 | -103.1437
20.1215 | -103.1299
20.1252 | -103.1194
20.1288 | -103.1081
20.1324 | -103.096
20.1361 | -103.0836
20.1397 | -103.0705
20.1433 | -103.0558
20.1469 | -103.0425
20.1506 | -103.0292
20.1542 | -103.016
20.1578 | -103.0037
20.1615 | -102.9925
20.1651 | -102.9811
20.1687 | -102.972
20.1723 | -102.9629
20.176 | -102.9533
20.1796 | -102.9451
20.1832 | -102.9384
20.1868 | -102.9334
20.1905 | -102.9269
20.1941 | -102.9192
20.1977 | -102.9137
20.2014 | -102.9078
20.205 | -102.9037
20.2086 | -102.8984
20.2122 | -102.8935
20.2159 | -102.8898
20.2195 | -102.8854
20.2231 | -102.8819
20.2268 | -102.8816
20.2304 | -102.8808
20.234 | -102.8789
20.2376 | -102.8753
20.2413 | -102.8708
20.2449 | -102.8675
20.2485 | -102.8667
20.2522 | -102.8661
20.2558 | -102.8637
20.2594 | -102.8607
20.263 | -102.8585
20.2667 | -102.8574
20.2703 | -102.8572
20.2739 | -102.8571
20.2776 | -102.8569
20.2812 | -102.8561
20.2848 | -102.8555
20.2884 | -102.8569
20.2921 | -102.8592
20.2957 | -102.8608
20.2993 | -102.8621
20.3029 | -102.8638
20.3066 | -102.8624
20.3102 | -102.862
20.3138 | -102.8646
20.3175 | -102.8664
20.3211 | -102.8689
20.3247 | -102.8699
20.3283 | -102.8708
20.332 | -102.8724
20.3356 | -102.8745
20.3392 | -102.8771
20.3429 | -102.8777
20.3465 | -102.8787
20.3501 | -102.8802
20.3537 | -102.8784
20.3574 | -102.8782
20.361 | -102.8819
20.3646 | -102.8819
20.3683 | -102.8828
20.3719 | -102.8851
20.3755 | -102.8848
20.3791 | -102.8841
20.3828 | -102.8844
20.3864 | -102.8858
20.39 | -102.8877
20.3937 | -102.889
20.3973 | -102.8861
20.4009 | -102.8835
20.4045 | -102.8845
20.4082 | -102.884
20.4118 | -102.8808
20.4154 | -102.8758
20.419 | -102.8717
20.4227 | -102.8698
20.4263 | -102.8697
20.4299 | -102.8687
20.4336 | -102.8687
20.4372 | -102.8661
20.4408 | -102.8633
20.4444 | -102.8616
20.4481 | -102.8577
20.4517 | -102.854
20.4553 | -102.8504
20.459 | -102.8458
20.4626 | -102.8435
20.4662 | -102.841
20.4698 | -102.8369
20.4735 | -102.8332
20.4771 | -102.8278
20.4807 | -102.8229
20.4844 | -102.8209
20.488 | -102.8163
20.4916 | -102.8133
20.4952 | -102.8108
20.4989 | -102.8079
20.5025 | -102.8053
20.5061 | -102.8029
20.5098 | -102.801
20.5134 | -102.7968
20.517 | -102.7908
20.5206 | -102.786
20.5243 | -102.7835
20.5279 | -102.7807
20.5315 | -102.7783
20.5351 | -102.7757
20.5388 | -102.7726
20.5424 | -102.7703
20.546 | -102.7676
20.5497 | -102.7651
20.5533 | -102.7622
20.5569 | -102.7588
20.5605 | -102.7556
20.5642 | -102.7529
20.5678 | -102.749
20.5714 | -102.7456
20.5751 | -102.744
20.5787 | -102.7435
20.5823 | -102.7434
20.5859 | -102.7452
20.5896 | -102.7443
20.5932 | -102.7429
20.5968 | -102.7442
20.6005 | -102.7459
20.6041 | -102.7451
20.6077 | -102.7433
20.6113 | -102.7399
20.615 | -102.7407
20.6186 | -102.7419
20.6222 | -102.7427
20.6258 | -102.7423
20.6295 | -102.7418
20.6331 | -102.7441
20.6367 | -102.7435
20.6404 | -102.7414
20.644 | -102.74
20.6476 | -102.7386
20.6512 | -102.7364
20.6549 | -102.7354
20.6585 | -102.734
20.6621 | -102.7321
20.6658 | -102.7288
20.6694 | -102.7246
20.673 | -102.7204
20.6766 | -102.7156
20.6803 | -102.712
20.6839 | -102.7104
20.6875 | -102.7097
20.6912 | -102.7088
20.6948 | -102.706
20.6984 | -102.7025
20.702 | -102.6999
20.7057 | -102.6976
20.7093 | -102.6926
20.7129 | -102.6869
20.7166 | -102.6804
20.7202 | -102.6752
20.7238 | -102.6721
20.7274 | -102.6665
20.7311 | -102.66
20.7347 | -102.6563
20.7383 | -102.6512
20.7419 | -102.6458
20.7456 | -102.6417
20.7492 | -102.6389
20.7528 | -102.6335
20.7565 | -102.6242
20.7601 | -102.6184
20.7637 | -102.6121
20.7673 | -102.6029
20.771 | -102.5951
20.7746 | -102.5892
20.7782 | -102.5804
20.7819 | -102.5701
20.7855 | -102.5593
20.7891 | -102.5515
20.7927 | -102.5438
20.7964 | -102.5348
20.8 | -102.5256
20.8036 | -102.5157
20.8073 | -102.5066
20.8109 | -102.4997
20.8145 | -102.4949
20.8181 | -102.4888
20.8218 | -102.4832
20.8254 | -102.4777
20.829 | -102.471
20.8327 | -102.4655
20.8363 | -102.4616
20.8399 | -102.4604
20.8435 | -102.4607
20.8472 | -102.4609
20.8508 | -102.4611
20.8544 | -102.4621
20.858 | -102.4627
20.8617 | -102.4645
20.8653 | -102.4674
20.8689 | -102.4711
20.8726 | -102.4739
20.8762 | -102.4786
20.8798 | -102.4837
20.8834 | -102.4915
20.8871 | -102.5002
20.8907 | -102.5089
20.8943 | -102.5184
20.898 | -102.5298
20.9016 | -102.542
20.9052 | -102.5549
20.9088 | -102.5645
20.9125 | -102.573
20.9161 | -102.5835
20.9197 | -102.596
20.9234 | -102.6111
20.927 | -102.6236
20.9306 | -102.6364
20.9342 | -102.6498
20.9379 | -102.6655
20.9415 | -102.6807
20.9451 | -102.6949
20.9488 | -102.7076
20.9524 | -102.7219
20.956 | -102.7363
20.9596 | -102.7473
20.9633 | -102.7608
20.9669 | -102.7752
20.9705 | -102.787
20.9742 | -102.7993
20.9778 | -102.8106
20.9814 | -102.8212
20.985 | -102.833
20.9887 | -102.8438
20.9923 | -102.8551
20.9959 | -102.8654
20.9995 | -102.8753
21.0032 | -102.8862
21.0068 | -102.8994
21.0104 | -102.9122
21.0141 | -102.9236
21.0177 | -102.9325
21.0213 | -102.9404
21.0249 | -102.9494
21.0286 | -102.9572
21.0322 | -102.9619
21.0358 | -102.9683
21.0395 | -102.9764
21.0431 | -102.9835
21.0467 | -102.9877
21.0503 | -102.9935
21.054 | -103.0023
21.0576 | -103.0088
21.0612 | -103.0172
21.0649 | -103.0278
21.0685 | -103.0362
21.0721 | -103.0418
21.0757 | -103.0498
21.0794 | -103.0594
21.083 | -103.0685
21.0866 | -103.074
21.0902 | -103.0824
21.0939 | -103.0917
21.0975 | -103.0999
21.1011 | -103.109
21.1048 | -103.1194
21.1084 | -103.1295
21.112 | -103.1345
21.1156 | -103.1412
21.1193 | -103.1505
21.1229 | -103.1601
21.1265 | -103.169
21.1302 | -103.1756
21.1338 | -103.1826
21.1374 | -103.1898
21.141 | -103.1969
21.1447 | -103.2039
21.1483 | -103.2099
21.1519 | -103.217
21.1556 | -103.224
21.1592 | -103.2292
21.1628 | -103.2355
21.1664 | -103.2411
21.1701 | -103.2461
21.1737 | -103.2502
21.1773 | -103.2547
21.181 | -103.2597
21.1846 | -103.2636
21.1882 | -103.2667
21.1918 | -103.2683
21.1955 | -103.2702
21.1991 | -103.2708
21.2027 | -103.2734
21.2063 | -103.2745
21.21 | -103.2766
21.2136 | -103.2788
21.2172 | -103.2804
21.2209 | -103.2802
21.2245 | -103.2806
21.2281 | -103.2828
21.2317 | -103.2851
21.2354 | -103.287
21.239 | -103.2876
21.2426 | -103.2897
21.2463 | -103.2903
21.2499 | -103.2899
21.2535 | -103.2894
21.2571 | -103.2903
21.2608 | -103.2927
21.2644 | -103.2921
21.268 | -103.2903
21.2717 | -103.2901
21.2753 | -103.2901
21.2789 | -103.2894
21.2825 | -103.287
21.2862 | -103.2862
21.2898 | -103.2862
21.2934 | -103.2864
21.2971 | -103.287
21.3007 | -103.2877
21.3043 | -103.2888
21.3079 | -103.2889
21.3116 | -103.2875
21.3152 | -103.2835
21.3188 | -103.2797
21.3224 | -103.2799
21.3261 | -103.2794
21.3297 | -103.2814
21.3333 | -103.2823
21.337 | -103.2778
21.3406 | -103.2744
21.3442 | -103.2705
21.3478 | -103.2693
21.3515 | -103.2664
21.3551 | -103.2629
21.3587 | -103.2593
21.3624 | -103.2544
21.366 | -103.2513
21.3696 | -103.249
21.3732 | -103.2467
21.3769 | -103.2425
21.3805 | -103.2385
21.3841 | -103.2338
21.3878 | -103.2284
21.3914 | -103.2226
21.395 | -103.2145
21.3986 | -103.2073
21.4023 | -103.1985
21.4059 | -103.1903
21.4095 | -103.1815
21.4132 | -103.1746
21.4168 | -103.165
21.4204 | -103.1526
21.424 | -103.1412
21.4277 | -103.1318
21.4313 | -103.1219
21.4349 | -103.1083
21.4385 | -103.0942
21.4422 | -103.0814
21.4458 | -103.0693
21.4494 | -103.0582
21.4531 | -103.0469
21.4567 | -103.0351
21.4603 | -103.0236
21.4639 | -103.0142
21.4676 | -103.0027
21.4712 | -102.9891
21.4748 | -102.975
21.4785 | -102.9627
21.4821 | -102.9499
21.4857 | -102.9363
21.4893 | -102.9244
21.493 | -102.9138
21.4966 | -102.9033
21.5002 | -102.8918
21.5039 | -102.8801
21.5075 | -102.8674
21.5111 | -102.8569
21.5147 | -102.8453
21.5184 | -102.8345
21.522 | -102.8208
21.5256 | -102.8065
21.5293 | -102.7955
21.5329 | -102.784
21.5365 | -102.7716
21.5401 | -102.7587
21.5438 | -102.7461
21.5474 | -102.7321
21.551 | -102.7198
21.5546 | -102.7072
21.5583 | -102.6923
21.5619 | -102.6797
21.5655 | -102.6682
21.5692 | -102.6574
21.5728 | -102.6466
21.5764 | -102.6343
21.58 | -102.6208
21.5837 | -102.6076
21.5873 | -102.5944
21.5909 | -102.5823
21.5946 | -102.571
21.5982 | -102.5572
21.6018 | -102.5443
21.6054 | -102.5314
21.6091 | -102.516
21.6127 | -102.5007
21.6163 | -102.489
21.62 | -102.4794
21.6236 | -102.4684
21.6272 | -102.4537
21.6308 | -102.4415
21.6345 | -102.4287
21.6381 | -102.414
21.6417 | -102.4013
21.6454 | -102.3857
21.649 | -102.3709
21.6526 | -102.3568
21.6562 | -102.3444
21.6599 | -102.3307
21.6635 | -102.3158
21.6671 | -102.3019
21.6707 | -102.2879
21.6744 | -102.2763
21.678 | -102.2643
21.6816 | -102.2518
21.6853 | -102.238
21.6889 | -102.2261
21.6925 | -102.215
21.6961 | -102.2047
21.6998 | -102.1961
21.7034 | -102.1878
21.707 | -102.1809
21.7107 | -102.1727
21.7143 | -102.1645
21.7179 | -102.1562
21.7215 | -102.1495
21.7252 | -102.1427
21.7288 | -102.1341
21.7324 | -102.1244
21.7361 | -102.115
21.7397 | -102.1095
21.7433 | -102.1048
21.7469 | -102.0997
21.7506 | -102.0928
21.7542 | -102.0836
21.7578 | -102.0768
21.7615 | -102.0719
21.7651 | -102.067
21.7687 | -102.0625
21.7723 | -102.0572
21.776 | -102.0512
21.7796 | -102.0471
21.7832 | -102.0451
21.7868 | -102.0413
21.7905 | -102.0374
21.7941 | -102.0357
21.7977 | -102.0314
21.8014 | -102.0263
21.805 | -102.0218
21.8086 | -102.0169
21.8122 | -102.0148
21.8159 | -102.0098
21.8195 | -102.0057
21.8231 | -102.0024
21.8268 | -101.9963
21.8304 | -101.9906
21.834 | -101.9852
21.8376 | -101.9776
21.8413 | -101.9701
21.8449 | -101.9649
21.8485 | -101.9603
21.8522 | -101.9557
21.8558 | -101.9497
21.8594 | -101.9437
21.863 | -101.9395
21.8667 | -101.9342
21.8703 | -101.9259
21.8739 | -101.9179
21.8776 | -101.9091
21.8812 | -101.9016
21.8848 | -101.895
21.8884 | -101.8881
21.8921 | -101.8801
21.8957 | -101.8721
21.8993 | -101.8622
21.9029 | -101.8519
21.9066 | -101.8436
21.9102 | -101.8356
21.9138 | -101.8279
21.9175 | -101.8215
21.9211 | -101.816
21.9247 | -101.8084
21.9283 | -101.8004
21.932 | -101.7931
21.9356 | -101.7883
21.9392 | -101.7821
21.9429 | -101.7776
21.9465 | -101.776
21.9501 | -101.7749
21.9537 | -101.7724
21.9574 | -101.7687
21.961 | -101.7651
21.9646 | -101.763
21.9683 | -101.7608
21.9719 | -101.7583
21.9755 | -101.7556
21.9791 | -101.7546
21.9828 | -101.7577
21.9864 | -101.7562
21.99 | -101.7576
21.9937 | -101.7586
21.9973 | -101.7606
22.0009 | -101.762
22.0045 | -101.762
22.0082 | -101.7617
22.0118 | -101.7634
22.0154 | -101.7658
22.019 | -101.767
22.0227 | -101.7679
22.0263 | -101.7685
22.0299 | -101.7697
22.0336 | -101.7708
22.0372 | -101.7721
22.0408 | -101.7733
22.0444 | -101.7747
22.0481 | -101.7765
22.0517 | -101.7762
22.0553 | -101.7741
22.059 | -101.7716
22.0626 | -101.7703
22.0662 | -101.7697
22.0698 | -101.766
22.0735 | -101.7614
22.0771 | -101.7566
22.0807 | -101.7509
22.0844 | -101.7455
22.088 | -101.7398
22.0916 | -101.7326
22.0952 | -101.7243
22.0989 | -101.7174
22.1025 | -101.7078
22.1061 | -101.6946
22.1098 | -101.6858
22.1134 | -101.6762
22.117 | -101.6669
22.1206 | -101.6594
22.1243 | -101.6508
22.1279 | -101.643
22.1315 | -101.6344
22.1351 | -101.6244
22.1388 | -101.614
22.1424 | -101.603
22.146 | -101.5914
22.1497 | -101.5819
22.1533 | -101.5715
22.1569 | -101.5575
22.1605 | -101.5452
22.1642 | -101.5334
22.1678 | -101.5228
22.1714 | -101.5126
22.1751 | -101.5014
22.1787 | -101.4905
22.1823 | -101.4816
22.1859 | -101.4704
22.1896 | -101.4591
22.1932 | -101.4493
22.1968 | -101.4382
22.2005 | -101.4285
22.2041 | -101.4195
22.2077 | -101.4112
22.2113 | -101.4036
22.215 | -101.3953
22.2186 | -101.3868
22.2222 | -101.3791
22.2259 | -101.3705
22.2295 | -101.3621
22.2331 | -101.3539
22.2367 | -101.344
22.2404 | -101.3348
22.244 | -101.327
22.2476 | -101.3175
22.2512 | -101.3098
22.2549 | -101.3004
22.2585 | -101.2908
22.2621 | -101.2821
22.2658 | -101.2748
22.2694 | -101.2664
22.273 | -101.2583
22.2766 | -101.2491
22.2803 | -101.2411
22.2839 | -101.2329
22.2875 | -101.2234
22.2912 | -101.2124
22.2948 | -101.2031
22.2984 | -101.1929
22.302 | -101.1799
22.3057 | -101.1641
22.3093 | -101.1505
22.3129 | -101.1391
22.3166 | -101.1277
22.3202 | -101.1165
22.3238 | -101.1041
22.3274 | -101.089
22.3311 | -101.0744
22.3347 | -101.0623
22.3383 | -101.051
22.342 | -101.0395
22.3456 | -101.0268
22.3492 | -101.0132
22.3528 | -100.9998
22.3565 | -100.9864
22.3601 | -100.973
22.3637 | -100.9612
22.3673 | -100.9501
22.371 | -100.9405
22.3746 | -100.9309
22.3782 | -100.9201
22.3819 | -100.9104
22.3855 | -100.9002
22.3891 | -100.8886
22.3927 | -100.8783
22.3964 | -100.8706
22.4 | -100.8619
22.4036 | -100.8542
22.4073 | -100.8487
22.4109 | -100.842
22.4145 | -100.8348
22.4181 | -100.8305
22.4218 | -100.827
22.4254 | -100.8212
22.429 | -100.8177
22.4327 | -100.8136
22.4363 | -100.8077
22.4399 | -100.8014
22.4435 | -100.7959
22.4472 | -100.7885
22.4508 | -100.7831
22.4544 | -100.7806
22.4581 | -100.7782
22.4617 | -100.7763
22.4653 | -100.7719
22.4689 | -100.7647
22.4726 | -100.7572
22.4762 | -100.7492
22.4798 | -100.74
22.4834 | -100.7321
22.4871 | -100.7231
22.4907 | -100.7137
22.4943 | -100.7048
22.498 | -100.6948
22.5016 | -100.6835
22.5052 | -100.6622
22.5088 | -100.6507
22.5125 | -100.6343
22.5161 | -100.6179
22.5197 | -100.6007
22.5234 | -100.5823
22.527 | -100.5623
22.5306 | -100.5395
22.5342 | -100.5173
22.5379 | -100.4971
22.5415 | -100.4747
22.5451 | -100.4514
22.5488 | -100.4281
22.5524 | -100.4033
22.556 | -100.3798
22.5596 | -100.3546
22.5633 | -100.3286
22.5669 | -100.3044
22.5705 | -100.28
22.5741 | -100.2553
22.5778 | -100.2291
22.5814 | -100.2039
22.585 | -100.1806
22.5887 | -100.1573
22.5923 | -100.1331
22.5959 | -100.1094
22.5995 | -100.0856
22.6032 | -100.0629
22.6068 | -100.042
22.6104 | -100.0222
22.6141 | -100.0046
22.6177 | -99.9869
22.6213 | -99.9713
22.6249 | -99.9593
22.6286 | -99.9474
22.6322 | -99.9358
22.6358 | -99.924
22.6395 | -99.9139
22.6431 | -99.9065
22.6467 | -99.9
22.6503 | -99.8926
22.654 | -99.8872
22.6576 | -99.883
22.6612 | -99.8815
22.6649 | -99.8827
22.6685 | -99.8851
22.6721 | -99.8879
22.6757 | -99.8918
22.6794 | -99.8977
22.683 | -99.902
22.6866 | -99.9067
22.6902 | -99.9126
22.6939 | -99.9197
22.6975 | -99.9282
22.7011 | -99.938
22.7048 | -99.949
22.7084 | -99.962
22.712 | -99.9733
22.7156 | -99.9822
22.7193 | -99.9938
22.7229 | -100.0043
22.7265 | -100.015
22.7302 | -100.0258
22.7338 | -100.0359
22.7374 | -100.0463
22.741 | -100.0572
22.7447 | -100.0666
22.7483 | -100.0733
22.7519 | -100.0813
22.7556 | -100.0902
22.7592 | -100.0957
22.7628 | -100.1004
22.7664 | -100.1057
22.7701 | -100.1097
22.7737 | -100.1113
22.7773 | -100.1131
22.781 | -100.1144
22.7846 | -100.1163
22.7882 | -100.1194
22.7918 | -100.1218
22.7955 | -100.121
22.7991 | -100.1196
22.8027 | -100.1146
22.8063 | -100.1082
22.81 | -100.1026
22.8136 | -100.095
22.8172 | -100.0844
22.8209 | -100.0748
22.8245 | -100.0665
22.8281 | -100.0581
22.8317 | -100.0479
22.8354 | -100.0376
22.839 | -100.028
22.8426 | -100.0187
22.8463 | -100.01
22.8499 | -100.0018
22.8535 | -99.9937
22.8571 | -99.9828
22.8608 | -99.9672
22.8644 | -99.9528
22.868 | -99.9386
22.8717 | -99.9258
22.8753 | -99.9146
22.8789 | -99.9014
22.8825 | -99.8896
22.8862 | -99.8778
22.8898 | -99.8654
22.8934 | -99.8541
22.8971 | -99.8437
22.9007 | -99.8363
22.9043 | -99.8277
22.9079 | -99.8192
22.9116 | -99.8124
22.9152 | -99.8056
22.9188 | -99.7988
22.9224 | -99.7901
22.9261 | -99.7835
22.9297 | -99.7793
22.9333 | -99.7732
22.937 | -99.7634
22.9406 | -99.7583
22.9442 | -99.7548
22.9478 | -99.7511
22.9515 | -99.7464
22.9551 | -99.7417
22.9587 | -99.7393
22.9624 | -99.7364
22.966 | -99.7335
22.9696 | -99.7323
22.9732 | -99.7333
22.9769 | -99.7329
22.9805 | -99.7334
22.9841 | -99.7332
22.9878 | -99.7338
22.9914 | -99.736
22.995 | -99.7395
22.9986 | -99.7422
23.0023 | -99.7436
23.0059 | -99.7463
23.0095 | -99.7505
23.0132 | -99.7557
23.0168 | -99.7601
23.0204 | -99.7654
23.024 | -99.7702
23.0277 | -99.7803
23.0313 | -99.7862
23.0349 | -99.792
23.0385 | -99.797
23.0422 | -99.8033
23.0458 | -99.8089
23.0494 | -99.8157
23.0531 | -99.8208
23.0567 | -99.8279
23.0603 | -99.8355
23.0639 | -99.8418
23.0676 | -99.848
23.0712 | -99.8551
23.0748 | -99.8621
23.0785 | -99.8677
23.0821 | -99.8729
23.0857 | -99.8801
23.0893 | -99.8844
23.093 | -99.8899
23.0966 | -99.8934
23.1002 | -99.8957
23.1039 | -99.8996
23.1075 | -99.9021
23.1111 | -99.9034
23.1147 | -99.9043
23.1184 | -99.9048
23.122 | -99.9068
23.1256 | -99.9086
23.1293 | -99.9117
23.1329 | -99.9127
23.1365 | -99.9092
23.1401 | -99.9065
23.1438 | -99.9032
23.1474 | -99.9008
23.151 | -99.8942
23.1546 | -99.8871
23.1583 | -99.8813
23.1619 | -99.8761
23.1655 | -99.8697
23.1692 | -99.8627
23.1728 | -99.8528
23.1764 | -99.8436
23.18 | -99.8344
23.1837 | -99.8283
23.1873 | -99.8231
23.1909 | -99.8157
23.1946 | -99.8091
23.1982 | -99.8005
23.2018 | -99.7918
23.2054 | -99.784
23.2091 | -99.777
23.2127 | -99.7691
23.2163 | -99.7609
23.22 | -99.7536
23.2236 | -99.7486
23.2272 | -99.7434
23.2308 | -99.7383
23.2345 | -99.7333
23.2381 | -99.7311
23.2417 | -99.7309
23.2454 | -99.7307
23.249 | -99.7326
23.2526 | -99.735
23.2562 | -99.736
23.2599 | -99.7389
23.2635 | -99.7439
23.2671 | -99.7486
23.2707 | -99.754
23.2744 | -99.7586
23.278 | -99.7642
23.2816 | -99.7711
23.2853 | -99.7791
23.2889 | -99.7888
23.2925 | -99.7991
23.2961 | -99.8084
23.2998 | -99.818
23.3034 | -99.8279
23.307 | -99.8388
23.3107 | -99.8498
23.3143 | -99.8601
23.3179 | -99.8709
23.3215 | -99.8814
23.3252 | -99.8934
23.3288 | -99.9056
23.3324 | -99.9184
23.3361 | -99.9317
23.3397 | -99.9432
23.3433 | -99.9554
23.3469 | -99.9688
23.3506 | -99.9805
23.3542 | -99.9901
23.3578 | -99.9989
23.3615 | -100.0079
23.3651 | -100.0182
23.3687 | -100.0311
23.3723 | -100.0426
23.376 | -100.0542
23.3796 | -100.0652
23.3832 | -100.0754
23.3868 | -100.0851
23.3905 | -100.0946
23.3941 | -100.1041
23.3977 | -100.1142
23.4014 | -100.1249
23.405 | -100.1344
23.4086 | -100.1426
23.4122 | -100.151
23.4159 | -100.16
23.4195 | -100.1708
23.4231 | -100.1785
23.4268 | -100.1877
23.4304 | -100.1983
23.434 | -100.206
23.4376 | -100.2129
23.4413 | -100.2213
23.4449 | -100.2307
23.4485 | -100.2409
23.4522 | -100.2503
23.4558 | -100.2594
23.4594 | -100.2674
23.463 | -100.2772
23.4667 | -100.2901
23.4703 | -100.3009
23.4739 | -100.3134
23.4776 | -100.324
23.4812 | -100.3366
23.4848 | -100.3467
23.4884 | -100.3562
23.4921 | -100.3669
23.4957 | -100.3765
23.4993 | -100.3858
23.5029 | -100.3969
23.5066 | -100.407
23.5102 | -100.4185
23.5138 | -100.4267
23.5175 | -100.4375
23.5211 | -100.4503
23.5247 | -100.4619
23.5283 | -100.4728
23.532 | -100.4824
23.5356 | -100.4936
23.5392 | -100.5031
23.5429 | -100.515
23.5465 | -100.5272
23.5501 | -100.5375
23.5537 | -100.5473
23.5574 | -100.5546
23.561 | -100.5615
23.5646 | -100.5716
23.5683 | -100.5818
23.5719 | -100.5882
23.5755 | -100.5968
23.5791 | -100.6082
23.5828 | -100.6204
23.5864 | -100.6318
23.59 | -100.6435
23.5937 | -100.6556
23.5973 | -100.6688
23.6009 | -100.6795
23.6045 | -100.6901
23.6082 | -100.7025
23.6118 | -100.7165
23.6154 | -100.7288
23.619 | -100.7381
23.6227 | -100.7471
23.6263 | -100.7565
23.6299 | -100.7677
23.6336 | -100.7788
23.6372 | -100.7898
23.6408 | -100.7988
23.6444 | -100.8093
23.6481 | -100.8207
23.6517 | -100.8318
23.6553 | -100.8434
23.659 | -100.8537
23.6626 | -100.8659
23.6662 | -100.8779
23.6698 | -100.8864
23.6735 | -100.8949
23.6771 | -100.9031
23.6807 | -100.9114
23.6844 | -100.9222
23.688 | -100.9309
23.6916 | -100.9429
23.6952 | -100.9552
23.6989 | -100.965
23.7025 | -100.9724
23.7061 | -100.9808
23.7098 | -100.9887
23.7134 | -100.997
23.717 | -101.0066
23.7206 | -101.0142
23.7243 | -101.0219
23.7279 | -101.0301
23.7315 | -101.0384
23.7351 | -101.046
23.7388 | -101.0563
23.7424 | -101.0678
23.746 | -101.0773
23.7497 | -101.0843
23.7533 | -101.091
23.7569 | -101.0996
23.7605 | -101.1065
23.7642 | -101.1119
23.7678 | -101.1191
23.7714 | -101.1269
23.7751 | -101.1316
23.7787 | -101.1355
23.7823 | -101.1408
23.7859 | -101.1461
23.7896 | -101.1516
23.7932 | -101.1584
23.7968 | -101.1663
23.8005 | -101.1724
23.8041 | -101.1786
23.8077 | -101.1846
23.8113 | -101.1895
23.815 | -101.1939
23.8186 | -101.1969
23.8222 | -101.2013
23.8258 | -101.2036
23.8295 | -101.2073
23.8331 | -101.2115
23.8367 | -101.2163
23.8404 | -101.2205
23.844 | -101.224
23.8476 | -101.2283
23.8512 | -101.2313
23.8549 | -101.2333
23.8585 | -101.2344
23.8621 | -101.2368
23.8658 | -101.2402
23.8694 | -101.2425
23.873 | -101.2454
23.8766 | -101.2459
23.8803 | -101.2474
23.8839 | -101.2479
23.8875 | -101.2484
23.8912 | -101.2492
23.8948 | -101.2519
23.8984 | -101.2521
23.902 | -101.2497
23.9057 | -101.2482
23.9093 | -101.2469
23.9129 | -101.2436
23.9166 | -101.2392
23.9202 | -101.2356
23.9238 | -101.232
23.9274 | -101.2284
23.9311 | -101.225
23.9347 | -101.2177
23.9383 | -101.2111
23.9419 | -101.2042
23.9456 | -101.1982
23.9492 | -101.1918
23.9528 | -101.1857
23.9565 | -101.1813
23.9601 | -101.177
23.9637 | -101.171
23.9673 | -101.1647
23.971 | -101.1565
23.9746 | -101.1486
23.9782 | -101.1433
23.9819 | -101.1358
23.9855 | -101.1256
23.9891 | -101.1166
23.9927 | -101.1078
23.9964 | -101.1001
24 | -101.0915
24.0036 | -101.0828
24.0073 | -101.075
24.0109 | -101.068
24.0145 | -101.062
24.0181 | -101.0531
24.0218 | -101.0439
24.0254 | -101.0381
24.029 | -101.0294
24.0327 | -101.0185
24.0363 | -101.0098
24.0399 | -101.0034
24.0435 | -100.9968
24.0472 | -100.9909
24.0508 | -100.9867
24.0544 | -100.9814
24.0581 | -100.9759
24.0617 | -100.9707
24.0653 | -100.9664
24.0689 | -100.9614
24.0726 | -100.9558
24.0762 | -100.9544
24.0798 | -100.9501
24.0834 | -100.9472
24.0871 | -100.944
24.0907 | -100.9439
24.0943 | -100.9446
24.098 | -100.9435
24.1016 | -100.9436
24.1052 | -100.9437
24.1088 | -100.9425
24.1125 | -100.9406
24.1161 | -100.9421
24.1197 | -100.9435
24.1234 | -100.9439
24.127 | -100.9441
24.1306 | -100.9424
24.1342 | -100.9404
24.1379 | -100.9399
24.1415 | -100.9406
24.1451 | -100.9427
24.1488 | -100.9444
24.1524 | -100.9464
24.156 | -100.9483
24.1596 | -100.9516
24.1633 | -100.9543
24.1669 | -100.957
24.1705 | -100.9579
24.1742 | -100.9592
24.1778 | -100.9587
24.1814 | -100.9596
24.185 | -100.9602
24.1887 | -100.9594
24.1923 | -100.9583
24.1959 | -100.9584
24.1995 | -100.9589
24.2032 | -100.9589
24.2068 | -100.9585
24.2104 | -100.9576
24.2141 | -100.9556
24.2177 | -100.9534
24.2213 | -100.951
24.2249 | -100.9503
24.2286 | -100.9509
24.2322 | -100.9505
24.2358 | -100.9478
24.2395 | -100.9464
24.2431 | -100.9447
24.2467 | -100.9417
24.2503 | -100.9379
24.254 | -100.9309
24.2576 | -100.9249
24.2612 | -100.9209
24.2649 | -100.9148
24.2685 | -100.9067
24.2721 | -100.9001
24.2757 | -100.8973
24.2794 | -100.8932
24.283 | -100.888
24.2866 | -100.884
24.2902 | -100.8805
24.2939 | -100.8773
24.2975 | -100.8744
24.3011 | -100.8713
24.3048 | -100.8679
24.3084 | -100.864
24.312 | -100.8624
24.3156 | -100.8606
24.3193 | -100.8575
24.3229 | -100.8548
24.3265 | -100.8528
24.3302 | -100.8505
24.3338 | -100.8482
24.3374 | -100.8465
24.341 | -100.8437
24.3447 | -100.8413
24.3483 | -100.8396
24.3519 | -100.8393
24.3556 | -100.8381
24.3592 | -100.8375
24.3628 | -100.8367
24.3664 | -100.835
24.3701 | -100.8343
24.3737 | -100.8336
24.3773 | -100.8343
24.381 | -100.835
24.3846 | -100.8359
24.3882 | -100.8354
24.3918 | -100.834
24.3955 | -100.8352
24.3991 | -100.8382
24.4027 | -100.8422
24.4063 | -100.8452
24.41 | -100.8483
24.4136 | -100.8527
24.4172 | -100.8561
24.4209 | -100.8592
24.4245 | -100.8629
24.4281 | -100.8667
24.4317 | -100.8722
24.4354 | -100.8759
24.439 | -100.8791
24.4426 | -100.8852
24.4463 | -100.8922
24.4499 | -100.9006
24.4535 | -100.91
24.4571 | -100.92
24.4608 | -100.9283
24.4644 | -100.9375
24.468 | -100.9458
24.4717 | -100.9537
24.4753 | -100.9633
24.4789 | -100.9744
24.4825 | -100.9849
24.4862 | -100.9948
24.4898 | -101.005
24.4934 | -101.0144
24.4971 | -101.0225
24.5007 | -101.0317
24.5043 | -101.0415
24.5079 | -101.0491
24.5116 | -101.0571
24.5152 | -101.0665
24.5188 | -101.0748
24.5224 | -101.0852
24.5261 | -101.0926
24.5297 | -101.1009
24.5333 | -101.1103
24.537 | -101.1192
24.5406 | -101.1267
24.5442 | -101.1345
24.5478 | -101.1424
24.5515 | -101.1508
24.5551 | -101.1564
24.5587 | -101.1601
24.5624 | -101.1614
24.566 | -101.1621
24.5696 | -101.1648
24.5732 | -101.1645
24.5769 | -101.1605
24.5805 | -101.1531
24.5841 | -101.1445
24.5878 | -101.1301
24.5914 | -101.1076
24.595 | -101.0588
24.5986 | -100.9393
24.6023 | -100.8943
24.6059 | -100.8506
24.6095 | -100.8068
24.6132 | -100.7623
24.6168 | -100.7141
24.6204 | -100.6589
24.624 | -100.6025
24.6277 | -100.5464
24.6313 | -100.4851
24.6349 | -100.4192
24.6385 | -100.3523
24.6422 | -100.2833
24.6458 | -100.2107
24.6494 | -100.1356
24.6531 | -100.0579
24.6567 | -99.9783
24.6603 | -99.8964
24.6639 | -99.8137
24.6676 | -99.73
24.6712 | -99.6468
24.6748 | -99.563
24.6785 | -99.4773
24.6821 | -99.3923
24.6857 | -99.307
24.6893 | -99.2221
24.693 | -99.1389
24.6966 | -99.0577
24.7002 | -98.9777
24.7039 | -98.8999
24.7075 | -98.8252
24.7111 | -98.7535
24.7147 | -98.6847
24.7184 | -98.6185
24.722 | -98.5563
24.7256 | -98.4965
24.7293 | -98.4401
24.7329 | -98.3876
24.7365 | -98.3399
24.7401 | -98.2974
24.7438 | -98.2594
24.7474 | -98.2258
24.751 | -98.1973
24.7546 | -98.1747
24.7583 | -98.1571
24.7619 | -98.1442
24.7655 | -98.1372
24.7692 | -98.1364
24.7728 | -98.1401
24.7764 | -98.1487
24.78 | -98.1623
24.7837 | -98.1811
24.7873 | -98.2051
24.7909 | -98.2327
24.7946 | -98.2659
24.7982 | -98.3036
24.8018 | -98.3465
24.8054 | -98.3946
24.8091 | -98.4475
24.8127 | -98.5046
24.8163 | -98.5654
24.82 | -98.6311
24.8236 | -98.7012
24.8272 | -98.7739
24.8308 | -98.849
24.8345 | -98.9273
24.8381 | -99.0084
24.8417 | -99.0919
24.8454 | -99.1786
24.849 | -99.267
24.8526 | -99.3554
24.8562 | -99.4455
24.8599 | -99.5378
24.8635 | -99.6298
24.8671 | -99.7214
24.8707 | -99.8127
24.8744 | -99.9049
24.878 | -99.9956
24.8816 | -100.0844
24.8853 | -100.1717
24.8889 | -100.2576
24.8925 | -100.342
24.8961 | -100.4233
24.8998 | -100.5017
24.9034 | -100.5798
24.907 | -100.6567
24.9107 | -100.7308
24.9143 | -100.7991
24.9179 | -100.8616
24.9215 | -100.9236
24.9252 | -100.985
24.9288 | -101.0456
24.9324 | -101.1025
24.9361 | -101.1588
24.9397 | -101.2147
24.9433 | -101.2889
24.9469 | -101.3602
24.9506 | -101.4177
24.9542 | -101.4535
24.9578 | -101.4793
24.9615 | -101.4935
24.9651 | -101.505
24.9687 | -101.5177
24.9723 | -101.5286
24.976 | -101.5366
24.9796 | -101.5403
24.9832 | -101.5457
24.9868 | -101.5487
24.9905 | -101.5495
24.9941 | -101.5492
24.9977 | -101.5472
25.0014 | -101.5432
25.005 | -101.542
25.0086 | -101.5411
25.0122 | -101.5401
25.0159 | -101.5388
25.0195 | -101.5362
25.0231 | -101.5336
25.0268 | -101.532
25.0304 | -101.5318
25.034 | -101.5305
25.0376 | -101.5287
25.0413 | -101.5272
25.0449 | -101.5255
25.0485 | -101.5238
25.0522 | -101.5222
25.0558 | -101.5189
25.0594 | -101.5132
25.063 | -101.5078
25.0667 | -101.5045
25.0703 | -101.5026
25.0739 | -101.4991
25.0776 | -101.4967
25.0812 | -101.4957
25.0848 | -101.4951
25.0884 | -101.4943
25.0921 | -101.491
25.0957 | -101.4888
25.0993 | -101.4895
25.1029 | -101.4889
25.1066 | -101.4884
25.1102 | -101.4894
25.1138 | -101.492
25.1175 | -101.4918
25.1211 | -101.4956
25.1247 | -101.4971
25.1283 | -101.4994
25.132 | -101.5012
25.1356 | -101.5041
25.1392 | -101.5092
25.1429 | -101.5157
25.1465 | -101.523
25.1501 | -101.5285
25.1537 | -101.5345
25.1574 | -101.5385
25.161 | -101.5421
25.1646 | -101.5472
25.1683 | -101.5534
25.1719 | -101.5618
25.1755 | -101.5691
25.1791 | -101.5755
25.1828 | -101.5822
25.1864 | -101.5882
25.19 | -101.5952
25.1937 | -101.6025
25.1973 | -101.6114
25.2009 | -101.6196
25.2045 | -101.6264
25.2082 | -101.6325
25.2118 | -101.6396
25.2154 | -101.6447
25.219 | -101.6508
25.2227 | -101.6564
25.2263 | -101.6625
25.2299 | -101.6683
25.2336 | -101.6728
25.2372 | -101.6793
25.2408 | -101.6865
25.2444 | -101.691
25.2481 | -101.6936
25.2517 | -101.6973
25.2553 | -101.7015
25.259 | -101.707
25.2626 | -101.7115
25.2662 | -101.7171
25.2698 | -101.7225
25.2735 | -101.7266
25.2771 | -101.7279
25.2807 | -101.7302
25.2844 | -101.7334
25.288 | -101.736
25.2916 | -101.7375
25.2952 | -101.7404
25.2989 | -101.7435
25.3025 | -101.7477
25.3061 | -101.7518
25.3098 | -101.7534
25.3134 | -101.7526
25.317 | -101.7548
25.3206 | -101.7571
25.3243 | -101.7594
25.3279 | -101.7634
25.3315 | -101.7619
25.3351 | -101.7618
25.3388 | -101.7638
25.3424 | -101.7651
25.346 | -101.7647
25.3497 | -101.766
25.3533 | -101.7693
25.3569 | -101.774
25.3605 | -101.7764
25.3642 | -101.7781
25.3678 | -101.7802
25.3714 | -101.7802
25.3751 | -101.7813
25.3787 | -101.783
25.3823 | -101.785
25.3859 | -101.7871
25.3896 | -101.7904
25.3932 | -101.7926
25.3968 | -101.7968
25.4005 | -101.8018
25.4041 | -101.8071
25.4077 | -101.811
25.4113 | -101.8155
25.415 | -101.8225
25.4186 | -101.8298
25.4222 | -101.8366
25.4259 | -101.8435
25.4295 | -101.8506
25.4331 | -101.8563
25.4367 | -101.8628
25.4404 | -101.8719
25.444 | -101.8814
25.4476 | -101.8918
25.4512 | -101.8997
25.4549 | -101.9047
25.4585 | -101.9107
25.4621 | -101.918
25.4658 | -101.9253
25.4694 | -101.9314
25.473 | -101.9387
25.4766 | -101.9476
25.4803 | -101.955
25.4839 | -101.9619
25.4875 | -101.9681
25.4912 | -101.9749
25.4948 | -101.9825
25.4984 | -101.9879
25.502 | -101.9925
25.5057 | -101.9962
25.5093 | -102.0003
25.5129 | -102.0061
25.5166 | -102.0128
25.5202 | -102.0196
25.5238 | -102.0249
25.5274 | -102.0281
25.5311 | -102.0302
25.5347 | -102.032
25.5383 | -102.0352
25.5419 | -102.038
25.5456 | -102.0416
25.5492 | -102.0446
25.5528 | -102.0475
25.5565 | -102.0485
25.5601 | -102.0494
25.5637 | -102.0519
25.5673 | -102.055
25.571 | -102.0572
25.5746 | -102.0579
25.5782 | -102.0587
25.5819 | -102.0619
25.5855 | -102.0647
25.5891 | -102.0657
25.5927 | -102.0678
25.5964 | -102.0683
25.6 | -102.0687
25.6036 | -102.0712
25.6073 | -102.0732
25.6109 | -102.074
25.6145 | -102.0748
25.6181 | -102.0757
25.6218 | -102.0774
25.6254 | -102.079
25.629 | -102.0819
25.6327 | -102.0858
25.6363 | -102.0901
25.6399 | -102.0919
25.6435 | -102.0935
25.6472 | -102.0963
25.6508 | -102.0984
25.6544 | -102.1026
25.658 | -102.1064
25.6617 | -102.1102
25.6653 | -102.1124
25.6689 | -102.1155
25.6726 | -102.1204
25.6762 | -102.1268
25.6798 | -102.1322
25.6834 | -102.1361
25.6871 | -102.1398
25.6907 | -102.1456
25.6943 | -102.1529
25.698 | -102.1586
25.7016 | -102.163
25.7052 | -102.1684
25.7088 | -102.1746
25.7125 | -102.1801
25.7161 | -102.1866
25.7197 | -102.1922
25.7234 | -102.1976
25.727 | -102.2049
25.7306 | -102.2116
25.7342 | -102.22
25.7379 | -102.2287
25.7415 | -102.236
25.7451 | -102.2448
25.7488 | -102.2548
25.7524 | -102.2627
25.756 | -102.2724
25.7596 | -102.2815
25.7633 | -102.2896
25.7669 | -102.2973
25.7705 | -102.3072
25.7741 | -102.3194
25.7778 | -102.3302
25.7814 | -102.3404
25.785 | -102.3496
25.7887 | -102.3596
25.7923 | -102.3704
25.7959 | -102.3805
25.7995 | -102.3893
25.8032 | -102.398
25.8068 | -102.4078
25.8104 | -102.4186
25.8141 | -102.4304
25.8177 | -102.4418
25.8213 | -102.4531
25.8249 | -102.4626
25.8286 | -102.472
25.8322 | -102.4823
25.8358 | -102.4938
25.8395 | -102.5036
25.8431 | -102.5116
25.8467 | -102.5209
25.8503 | -102.5304
25.854 | -102.5393
25.8576 | -102.5454
25.8612 | -102.5522
25.8649 | -102.5599
25.8685 | -102.5693
25.8721 | -102.5776
25.8757 | -102.584
25.8794 | -102.5907
25.883 | -102.597
25.8866 | -102.6021
25.8902 | -102.609
25.8939 | -102.6167
25.8975 | -102.623
25.9011 | -102.6272
25.9048 | -102.6316
25.9084 | -102.6361
25.912 | -102.6405
25.9156 | -102.6439
25.9193 | -102.6492
25.9229 | -102.654
25.9265 | -102.6588
25.9302 | -102.664
25.9338 | -102.6694
25.9374 | -102.6764
25.941 | -102.6814
25.9447 | -102.6853
25.9483 | -102.6882
25.9519 | -102.6917
25.9556 | -102.695
25.9592 | -102.6982
25.9628 | -102.7015
25.9664 | -102.7049
25.9701 | -102.7096
25.9737 | -102.7149
25.9773 | -102.7198
25.981 | -102.7236
25.9846 | -102.7269
25.9882 | -102.7306
25.9918 | -102.7339
25.9955 | -102.7366
25.9991 | -102.7401
26.0027 | -102.7467
26.0063 | -102.7526
26.01 | -102.7589
26.0136 | -102.7648
26.0172 | -102.7709
26.0209 | -102.7745
26.0245 | -102.7786
26.0281 | -102.7832
26.0317 | -102.789
26.0354 | -102.7958
26.039 | -102.8032
26.0426 | -102.8095
26.0463 | -102.8166
26.0499 | -102.8253
26.0535 | -102.8338
26.0571 | -102.8417
26.0608 | -102.8502
26.0644 | -102.8599
26.068 | -102.8704
26.0717 | -102.8802
26.0753 | -102.8883
26.0789 | -102.896
26.0825 | -102.9038
26.0862 | -102.914
26.0898 | -102.9247
26.0934 | -102.933
26.0971 | -102.9408
26.1007 | -102.9503
26.1043 | -102.9609
26.1079 | -102.9701
26.1116 | -102.9777
26.1152 | -102.985
26.1188 | -102.9946
26.1224 | -103.0041
26.1261 | -103.0126
26.1297 | -103.0216
26.1333 | -103.0319
26.137 | -103.0388
26.1406 | -103.0433
26.1442 | -103.0493
26.1478 | -103.0551
26.1515 | -103.0598
26.1551 | -103.0645
26.1587 | -103.0699
26.1624 | -103.0776
26.166 | -103.0831
26.1696 | -103.0805
26.1732 | -103.0827
26.1769 | -103.0837
26.1805 | -103.0846
26.1841 | -103.0848
26.1878 | -103.0867
26.1914 | -103.0874
26.195 | -103.0858
26.1986 | -103.0837
26.2023 | -103.0831
26.2059 | -103.0819
26.2095 | -103.0802
26.2132 | -103.0793
26.2168 | -103.0773
26.2204 | -103.0752
26.224 | -103.0732
26.2277 | -103.0729
26.2313 | -103.0727
26.2349 | -103.0702
26.2385 | -103.0657
26.2422 | -103.0613
26.2458 | -103.0577
26.2494 | -103.0545
26.2531 | -103.051
26.2567 | -103.0488
26.2603 | -103.0449
26.2639 | -103.0402
26.2676 | -103.0378
26.2712 | -103.0362
26.2748 | -103.0349
26.2785 | -103.0332
26.2821 | -103.0308
26.2857 | -103.0278
26.2893 | -103.0242
26.293 | -103.0222
26.2966 | -103.0212
26.3002 | -103.0206
26.3039 | -103.0201
26.3075 | -103.0185
26.3111 | -103.0178
26.3147 | -103.0186
26.3184 | -103.02
26.322 | -103.0229
26.3256 | -103.0255
26.3293 | -103.0283
26.3329 | -103.0311
26.3365 | -103.0351
26.3401 | -103.0401
26.3438 | -103.0436
26.3474 | -103.0469
26.351 | -103.0506
26.3546 | -103.0562
26.3583 | -103.0627
26.3619 | -103.0672
26.3655 | -103.0712
26.3692 | -103.0771
26.3728 | -103.0822
26.3764 | -103.0875
26.38 | -103.0944
26.3837 | -103.1009
26.3873 | -103.1074
26.3909 | -103.1154
26.3946 | -103.1233
26.3982 | -103.1299
26.4018 | -103.1359
26.4054 | -103.1429
26.4091 | -103.1494
26.4127 | -103.1552
26.4163 | -103.1611
26.42 | -103.1659
26.4236 | -103.1704
26.4272 | -103.1751
26.4308 | -103.1811
26.4345 | -103.1864
26.4381 | -103.19
26.4417 | -103.1946
26.4454 | -103.1992
26.449 | -103.2054
26.4526 | -103.2094
26.4562 | -103.2119
26.4599 | -103.2136
26.4635 | -103.2157
26.4671 | -103.216
26.4707 | -103.2151
26.4744 | -103.2138
26.478 | -103.2128
26.4816 | -103.2116
26.4853 | -103.2101
26.4889 | -103.2068
26.4925 | -103.2036
26.4961 | -103.1995
26.4998 | -103.1922
26.5034 | -103.1842
26.507 | -103.1807
26.5107 | -103.1781
26.5143 | -103.1724
26.5179 | -103.1654
26.5215 | -103.159
26.5252 | -103.153
26.5288 | -103.1451
26.5324 | -103.1357
26.5361 | -103.1274
26.5397 | -103.1195
26.5433 | -103.1119
26.5469 | -103.1021
26.5506 | -103.0932
26.5542 | -103.0834
26.5578 | -103.0745
26.5615 | -103.0662
26.5651 | -103.0568
26.5687 | -103.0487
26.5723 | -103.0416
26.576 | -103.0315
26.5796 | -103.021
26.5832 | -103.0122
26.5868 | -103.0042
26.5905 | -102.9943
26.5941 | -102.9847
26.5977 | -102.9755
26.6014 | -102.967
26.605 | -102.9583
26.6086 | -102.9505
26.6122 | -102.9445
26.6159 | -102.9394
26.6195 | -102.9357
26.6231 | -102.9301
26.6268 | -102.9238
26.6304 | -102.918
26.634 | -102.9145
26.6376 | -102.9108
26.6413 | -102.9073
26.6449 | -102.9047
26.6485 | -102.9021
26.6522 | -102.9006
26.6558 | -102.8992
26.6594 | -102.8986
26.663 | -102.8981
26.6667 | -102.897
26.6703 | -102.8947
26.6739 | -102.8936
26.6776 | -102.893
26.6812 | -102.8933
26.6848 | -102.8934
26.6884 | -102.8918
26.6921 | -102.8898
26.6957 | -102.8877
26.6993 | -102.8862
26.7029 | -102.8872
26.7066 | -102.8883
26.7102 | -102.8882
26.7138 | -102.8863
26.7175 | -102.8848
26.7211 | -102.8839
26.7247 | -102.883
26.7283 | -102.8815
26.732 | -102.8788
26.7356 | -102.8772
26.7392 | -102.878
26.7429 | -102.8786
26.7465 | -102.8764
26.7501 | -102.873
26.7537 | -102.8699
26.7574 | -102.8659
26.761 | -102.8612
26.7646 | -102.8581
26.7683 | -102.8549
26.7719 | -102.8508
26.7755 | -102.8473
26.7791 | -102.8466
26.7828 | -102.8426
26.7864 | -102.837
26.79 | -102.8314
26.7937 | -102.8265
26.7973 | -102.8227
26.8009 | -102.8189
26.8045 | -102.8146
26.8082 | -102.8082
26.8118 | -102.7997
26.8154 | -102.7913
26.819 | -102.7832
26.8227 | -102.7785
26.8263 | -102.7728
26.8299 | -102.7668
26.8336 | -102.7604
26.8372 | -102.7536
26.8408 | -102.7466
26.8444 | -102.7393
26.8481 | -102.7318
26.8517 | -102.7229
26.8553 | -102.7159
26.859 | -102.7083
26.8626 | -102.6991
26.8662 | -102.6911
26.8698 | -102.6841
26.8735 | -102.6776
26.8771 | -102.6715
26.8807 | -102.6627
26.8844 | -102.6527
26.888 | -102.6428
26.8916 | -102.6337
26.8952 | -102.6244
26.8989 | -102.6146
26.9025 | -102.6047
26.9061 | -102.5952
26.9098 | -102.5865
26.9134 | -102.5788
26.917 | -102.569
26.9206 | -102.5564
26.9243 | -102.5463
26.9279 | -102.5376
26.9315 | -102.5286
26.9351 | -102.5163
26.9388 | -102.502
26.9424 | -102.4885
26.946 | -102.477
26.9497 | -102.4663
26.9533 | -102.4545
26.9569 | -102.4425
26.9605 | -102.4301
26.9642 | -102.4174
26.9678 | -102.4076
26.9714 | -102.3971
26.9751 | -102.3869
26.9787 | -102.3753
26.9823 | -102.3643
26.9859 | -102.3541
26.9896 | -102.3437
26.9932 | -102.3324
26.9968 | -102.3219
27.0005 | -102.3135
27.0041 | -102.3066
27.0077 | -102.2988
27.0113 | -102.2907
27.015 | -102.283
27.0186 | -102.2745
27.0222 | -102.2649
27.0258 | -102.2554
27.0295 | -102.2475
27.0331 | -102.2396
27.0367 | -102.2303
27.0404 | -102.2234
27.044 | -102.2169
27.0476 | -102.2117
27.0512 | -102.207
27.0549 | -102.2015
27.0585 | -102.1973
27.0621 | -102.1933
27.0658 | -102.1902
27.0694 | -102.1878
27.073 | -102.1879
27.0766 | -102.1877
27.0803 | -102.1854
27.0839 | -102.1831
27.0875 | -102.1811
27.0912 | -102.1815
27.0948 | -102.1825
27.0984 | -102.1828
27.102 | -102.1843
27.1057 | -102.187
27.1093 | -102.189
27.1129 | -102.192
27.1166 | -102.1935
27.1202 | -102.1951
27.1238 | -102.1983
27.1274 | -102.2034
27.1311 | -102.2098
27.1347 | -102.2171
27.1383 | -102.2226
27.142 | -102.2289
27.1456 | -102.2356
27.1492 | -102.2425
27.1528 | -102.2507
27.1565 | -102.2601
27.1601 | -102.267
27.1637 | -102.2757
27.1673 | -102.2858
27.171 | -102.2969
27.1746 | -102.3062
27.1782 | -102.3148
27.1819 | -102.325
27.1855 | -102.3348
27.1891 | -102.3451
27.1927 | -102.3558
27.1964 | -102.367
27.2 | -102.3783
27.2036 | -102.3869
27.2073 | -102.3959
27.2109 | -102.4056
27.2145 | -102.415
27.2181 | -102.4243
27.2218 | -102.4337
27.2254 | -102.4403
27.229 | -102.4475
27.2327 | -102.4547
27.2363 | -102.4636
27.2399 | -102.4704
27.2435 | -102.4755
27.2472 | -102.4833
27.2508 | -102.4904
27.2544 | -102.4963
27.2581 | -102.5015
27.2617 | -102.5078
27.2653 | -102.5146
27.2689 | -102.5189
27.2726 | -102.5238
27.2762 | -102.5291
27.2798 | -102.5335
27.2834 | -102.5367
27.2871 | -102.5391
27.2907 | -102.541
27.2943 | -102.5427
27.298 | -102.5438
27.3016 | -102.5464
27.3052 | -102.5509
27.3088 | -102.5528
27.3125 | -102.5517
27.3161 | -102.5516
27.3197 | -102.5547
27.3234 | -102.5572
27.327 | -102.5577
27.3306 | -102.5563
27.3342 | -102.5544
27.3379 | -102.5525
27.3415 | -102.5494
27.3451 | -102.5454
27.3488 | -102.5421
27.3524 | -102.5372
27.356 | -102.534
27.3596 | -102.5299
27.3633 | -102.5254
27.3669 | -102.5198
27.3705 | -102.5142
27.3742 | -102.5096
27.3778 | -102.5062
27.3814 | -102.4998
27.385 | -102.4939
27.3887 | -102.4878
27.3923 | -102.4812
27.3959 | -102.474
27.3995 | -102.4697
27.4032 | -102.4654
27.4068 | -102.4614
27.4104 | -102.4591
27.4141 | -102.4544
27.4177 | -102.451
27.4213 | -102.4478
27.4249 | -102.4418
27.4286 | -102.4369
27.4322 | -102.4308
27.4358 | -102.4236
27.4395 | -102.419
27.4431 | -102.4142
27.4467 | -102.4093
27.4503 | -102.4033
27.454 | -102.3986
27.4576 | -102.3924
27.4612 | -102.3866
27.4649 | -102.3815
27.4685 | -102.3743
27.4721 | -102.369
27.4757 | -102.3642
27.4794 | -102.3602
27.483 | -102.3572
27.4866 | -102.354
27.4902 | -102.3511
27.4939 | -102.3471
27.4975 | -102.3414
27.5011 | -102.3374
27.5048 | -102.3339
27.5084 | -102.3308
27.512 | -102.3271
27.5156 | -102.3234
27.5193 | -102.3204
27.5229 | -102.3191
27.5265 | -102.3163
27.5302 | -102.3121
27.5338 | -102.3088
27.5374 | -102.3063
27.541 | -102.3028
27.5447 | -102.3002
27.5483 | -102.2977
27.5519 | -102.2928
27.5556 | -102.2883
27.5592 | -102.2857
27.5628 | -102.2834
27.5664 | -102.2787
27.5701 | -102.2744
27.5737 | -102.271
27.5773 | -102.2686
27.581 | -102.2651
27.5846 | -102.2601
27.5882 | -102.2551
27.5918 | -102.2508
27.5955 | -102.2468
27.5991 | -102.2411
27.6027 | -102.2363
27.6063 | -102.2324
27.61 | -102.2285
27.6136 | -102.2272
27.6172 | -102.2265
27.6209 | -102.2262
27.6245 | -102.2242
27.6281 | -102.2217
27.6317 | -102.2201
27.6354 | -102.2165
27.639 | -102.2128
27.6426 | -102.2093
27.6463 | -102.2072
27.6499 | -102.2052
27.6535 | -102.2026
27.6571 | -102.2003
27.6608 | -102.1969
27.6644 | -102.1929
27.668 | -102.1894
27.6717 | -102.1871
27.6753 | -102.1846
27.6789 | -102.1819
27.6825 | -102.1789
27.6862 | -102.1749
27.6898 | -102.1706
27.6934 | -102.166
27.6971 | -102.1625
27.7007 | -102.1587
27.7043 | -102.1549
27.7079 | -102.1503
27.7116 | -102.1472
27.7152 | -102.1447
27.7188 | -102.1406
27.7224 | -102.1355
27.7261 | -102.1332
27.7297 | -102.1315
27.7333 | -102.1293
27.737 | -102.1248
27.7406 | -102.12
27.7442 | -102.1166
27.7478 | -102.1119
27.7515 | -102.1077
27.7551 | -102.1052
27.7587 | -102.1017
27.7624 | -102.097
27.766 | -102.0909
27.7696 | -102.0844
27.7732 | -102.0804
27.7769 | -102.0758
27.7805 | -102.0714
27.7841 | -102.0678
27.7878 | -102.0637
27.7914 | -102.0579
27.795 | -102.0523
27.7986 | -102.0474
27.8023 | -102.0425
27.8059 | -102.0388
27.8095 | -102.0351
27.8132 | -102.0326
27.8168 | -102.0292
27.8204 | -102.0249
27.824 | -102.022
27.8277 | -102.0192
27.8313 | -102.0157
27.8349 | -102.0117
27.8385 | -102.0102
27.8422 | -102.0079
27.8458 | -102.006
27.8494 | -102.0048
27.8531 | -102.0046
27.8567 | -102.004
27.8603 | -102.0041
27.8639 | -102.0037
27.8676 | -102.0016
27.8712 | -101.9994
27.8748 | -101.9984
27.8785 | -101.9976
27.8821 | -101.9952
27.8857 | -101.9931
27.8893 | -101.9926
27.893 | -101.9928
27.8966 | -101.9909
27.9002 | -101.9873
27.9039 | -101.9857
27.9075 | -101.9848
27.9111 | -101.9823
27.9147 | -101.9807
27.9184 | -101.9821
27.922 | -101.9826
27.9256 | -101.9804
27.9293 | -101.9777
27.9329 | -101.9757
27.9365 | -101.9735
27.9401 | -101.9716
27.9438 | -101.9725
27.9474 | -101.9715
27.951 | -101.9686
27.9546 | -101.9658
27.9583 | -101.9637
27.9619 | -101.9618
27.9655 | -101.9604
27.9692 | -101.9573
27.9728 | -101.9546
27.9764 | -101.9528
27.98 | -101.949
27.9837 | -101.9442
27.9873 | -101.9404
27.9909 | -101.9381
27.9946 | -101.9335
27.9982 | -101.9266
28.0018 | -101.9199