## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 227 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 231
Hood | 02-92B/P | 170
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1013.6 | hPa |  | PRC PRO 01
Humidity | 43.5 | %/Hr
Internal temperature of calibration bench | 23.21 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1017 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.2 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/ input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.9 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3106 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument | Type
Adjustment | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Metrix Multimeter | 34410A
| Oscilloscope | Tektronix TDS 3014B
| Power Supply | E36104A
| Acoustical and electrical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone | 4193-L-004
| Microphone power supply | 2829
| Meteo station | ALMEMO 2690
| Power Supply | E36104A
| Acoustical measurement system | PR 000488 -
| Acoustical generator | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer | Digitizer Hood
| Power Supply | E36104A
Self-noise measurement | Digitizer | TAURUS
| Power Supply | E36104A
| Vibration isolation | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 15/05/2018
| Ludovic LECLAIR | Measurement of electrical and acoustical transfer function Measurement of self-noise |  | 17/05/2018
Engineer | Anthony HUE | Approval |  | 18/05/2018
CEO | Robert GLEMOT | Approval |  | 18/05/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4911 | 57.5947
0.012209 | -36.0754 | 38.4734
0.018313 | -35.015 | 28.2912
0.024417 | -34.5551 | 21.9965
0.030521 | -34.433 | 18.2008
0.036626 | -34.2852 | 15.154
0.04273 | -34.1896 | 12.8114
0.048834 | -34.1439 | 11.2689
0.054938 | -34.0837 | 10.117
0.061043 | -34.0275 | 9.0533
0.067147 | -34.0164 | 8.2326
0.073251 | -34.0153 | 7.4708
0.079355 | -34.0298 | 6.8077
0.08546 | -34.0405 | 6.2189
0.091564 | -34.0345 | 5.778
0.097668 | -34.0172 | 5.5134
0.10377 | -33.9977 | 5.2557
0.10988 | -33.9824 | 4.9738
0.11598 | -33.9744 | 4.6827
0.12209 | -33.976 | 4.4041
0.12819 | -33.9844 | 4.1294
0.13429 | -33.9933 | 3.853
0.1404 | -33.9956 | 3.5845
0.1465 | -33.9969 | 3.3347
0.15261 | -33.9942 | 3.1188
0.15871 | -33.9885 | 2.9558
0.16482 | -33.9816 | 2.8471
0.17092 | -33.9798 | 2.798
0.17702 | -33.978 | 2.7197
0.18313 | -33.9773 | 2.6375
0.18923 | -33.9769 | 2.5295
0.19534 | -33.9756 | 2.4063
0.20144 | -33.9742 | 2.2696
0.20754 | -33.9714 | 2.1512
0.21365 | -33.9703 | 2.0387
0.21975 | -33.9697 | 1.941
0.22586 | -33.9705 | 1.853
0.23196 | -33.9727 | 1.779
0.23807 | -33.9763 | 1.7143
0.24417 | -33.9801 | 1.6552
0.25027 | -33.9838 | 1.5926
0.25638 | -33.9862 | 1.5309
0.26248 | -33.988 | 1.4657
0.26859 | -33.9892 | 1.4018
0.27469 | -33.9898 | 1.3331
0.2808 | -33.991 | 1.2622
0.2869 | -33.9919 | 1.1938
0.293 | -33.9929 | 1.1288
0.29911 | -33.9939 | 1.0699
0.30521 | -33.9948 | 1.012
0.31132 | -33.9953 | 0.97275
0.31742 | -33.9963 | 0.92971
0.32353 | -33.9964 | 0.89364
0.32963 | -33.997 | 0.86305
0.33573 | -33.9976 | 0.84042
0.34184 | -33.9986 | 0.8169
0.34794 | -34.0003 | 0.79155
0.35405 | -34.0021 | 0.76586
0.36015 | -34.0042 | 0.73891
0.36626 | -34.0059 | 0.71017
0.37236 | -34.0077 | 0.67332
0.37846 | -34.0096 | 0.62991
0.38457 | -34.0105 | 0.57509
0.39067 | -34.0114 | 0.52955
0.39678 | -34.012 | 0.47951
0.40288 | -34.0125 | 0.4277
0.40899 | -34.0124 | 0.37765
0.41509 | -34.0119 | 0.32876
0.42119 | -34.0119 | 0.28711
0.4273 | -34.0119 | 0.24262
0.4334 | -34.0118 | 0.19838
0.43951 | -34.0121 | 0.15686
0.44561 | -34.0129 | 0.11868
0.45172 | -34.0143 | 0.087989
0.45782 | -34.015 | 0.052354
0.46392 | -34.0161 | 0.017574
0.47003 | -34.0171 | -0.013738
0.47613 | -34.0179 | -0.046488
0.48224 | -34.0184 | -0.079895
0.48834 | -34.0177 | -0.11324
0.49445 | -34.0175 | -0.14713
0.50055 | -34.0171 | -0.17812
0.50665 | -34.0165 | -0.21105
0.51276 | -34.0157 | -0.24389
0.51886 | -34.0148 | -0.2766
0.52497 | -34.0138 | -0.30083
0.53107 | -34.0131 | -0.32958
0.53717 | -34.0126 | -0.35731
0.54328 | -34.0119 | -0.38036
0.54938 | -34.0118 | -0.40307
0.55549 | -34.0118 | -0.42437
0.56159 | -34.0128 | -0.43746
0.5677 | -34.0131 | -0.4532
0.5738 | -34.0141 | -0.47043
0.5799 | -34.0148 | -0.48645
0.58601 | -34.0158 | -0.49824
0.59211 | -34.0173 | -0.50651
0.59822 | -34.0185 | -0.51758
0.60432 | -34.0199 | -0.5309
0.61043 | -34.021 | -0.54067
0.61653 | -34.0221 | -0.55227
0.62263 | -34.0235 | -0.5631
0.62874 | -34.0241 | -0.57495
0.63484 | -34.0249 | -0.58566
0.64095 | -34.0253 | -0.59716
0.64705 | -34.0256 | -0.60825
0.65316 | -34.0258 | -0.62161
0.65926 | -34.0259 | -0.63789
0.66536 | -34.0249 | -0.65545
0.67147 | -34.0246 | -0.66942
0.67757 | -34.0243 | -0.68428
0.68368 | -34.0237 | -0.70315
0.68978 | -34.023 | -0.72051
0.69589 | -34.0222 | -0.73938
0.70199 | -34.0214 | -0.76142
0.70809 | -34.0206 | -0.78041
0.7142 | -34.02 | -0.80046
0.7203 | -34.0191 | -0.81943
0.72641 | -34.0185 | -0.84124
0.73251 | -34.018 | -0.86083
0.73862 | -34.0182 | -0.88186
0.74472 | -34.018 | -0.90141
0.75082 | -34.0178 | -0.92276
0.75693 | -34.0181 | -0.94385
0.76303 | -34.0186 | -0.96507
0.76914 | -34.0202 | -0.98479
0.77524 | -34.0211 | -1.0039
0.78135 | -34.0223 | -1.0234
0.78745 | -34.0235 | -1.0427
0.79355 | -34.0246 | -1.0618
0.79966 | -34.0262 | -1.0807
0.80576 | -34.0282 | -1.0997
0.81187 | -34.0297 | -1.1203
0.81797 | -34.0314 | -1.1407
0.82408 | -34.0329 | -1.1604
0.83018 | -34.0345 | -1.1801
0.83628 | -34.0362 | -1.1981
0.84239 | -34.0373 | -1.2178
0.84849 | -34.0386 | -1.2386
0.8546 | -34.0399 | -1.2578
0.8607 | -34.041 | -1.2777
0.86681 | -34.042 | -1.2961
0.87291 | -34.0429 | -1.3144
0.87901 | -34.0433 | -1.332
0.88512 | -34.0439 | -1.3497
0.89122 | -34.0446 | -1.3676
0.89733 | -34.045 | -1.3831
0.90343 | -34.0454 | -1.4004
0.90953 | -34.0455 | -1.4149
0.91564 | -34.0456 | -1.4307
0.92174 | -34.0458 | -1.4457
0.92785 | -34.0459 | -1.4603
0.93395 | -34.046 | -1.4744
0.94006 | -34.0462 | -1.4875
0.94616 | -34.0459 | -1.4998
0.95226 | -34.0461 | -1.5127
0.95837 | -34.0462 | -1.5267
0.96447 | -34.0464 | -1.5398
0.97058 | -34.0462 | -1.5534
0.97668 | -34.0461 | -1.5669
0.98279 | -34.0459 | -1.5799
0.98889 | -34.0457 | -1.5911
0.99499 | -34.0457 | -1.6046
1.0011 | -34.0456 | -1.6157
1.0072 | -34.0453 | -1.6277
1.0133 | -34.0451 | -1.6389
1.0194 | -34.0447 | -1.652
1.0255 | -34.0445 | -1.6658
1.0316 | -34.044 | -1.678
1.0377 | -34.0437 | -1.6916
1.0438 | -34.0432 | -1.7045
1.0499 | -34.0426 | -1.7196
1.056 | -34.042 | -1.7338
1.0621 | -34.0415 | -1.7481
1.0682 | -34.0411 | -1.7606
1.0743 | -34.0405 | -1.7758
1.0805 | -34.0401 | -1.7901
1.0866 | -34.0393 | -1.8049
1.0927 | -34.0386 | -1.8192
1.0988 | -34.0381 | -1.8339
1.1049 | -34.0374 | -1.848
1.111 | -34.0367 | -1.8631
1.1171 | -34.0362 | -1.8774
1.1232 | -34.0358 | -1.8922
1.1293 | -34.0354 | -1.9061
1.1354 | -34.0352 | -1.9208
1.1415 | -34.0349 | -1.936
1.1476 | -34.0344 | -1.9511
1.1537 | -34.0339 | -1.9655
1.1598 | -34.0341 | -1.9799
1.1659 | -34.0339 | -1.993
1.172 | -34.0337 | -2.0067
1.1781 | -34.0335 | -2.0205
1.1842 | -34.0334 | -2.0344
1.1903 | -34.0337 | -2.0455
1.1964 | -34.0337 | -2.0587
1.2025 | -34.0338 | -2.0724
1.2086 | -34.0341 | -2.0847
1.2147 | -34.0342 | -2.0971
1.2209 | -34.0345 | -2.1083
1.227 | -34.0352 | -2.1192
1.2331 | -34.0356 | -2.1311
1.2392 | -34.0359 | -2.143
1.2453 | -34.0363 | -2.1541
1.2514 | -34.0367 | -2.1642
1.2575 | -34.0373 | -2.1753
1.2636 | -34.038 | -2.1846
1.2697 | -34.0386 | -2.1942
1.2758 | -34.0392 | -2.2054
1.2819 | -34.0399 | -2.2156
1.288 | -34.0405 | -2.2265
1.2941 | -34.0412 | -2.2354
1.3002 | -34.0418 | -2.244
1.3063 | -34.0423 | -2.2534
1.3124 | -34.0429 | -2.2628
1.3185 | -34.0436 | -2.2722
1.3246 | -34.0442 | -2.2817
1.3307 | -34.045 | -2.2907
1.3368 | -34.0457 | -2.2998
1.3429 | -34.0463 | -2.3091
1.349 | -34.0468 | -2.3188
1.3551 | -34.0474 | -2.3285
1.3613 | -34.0479 | -2.3372
1.3674 | -34.0485 | -2.3475
1.3735 | -34.0489 | -2.3575
1.3796 | -34.0493 | -2.3669
1.3857 | -34.0496 | -2.3765
1.3918 | -34.0501 | -2.3868
1.3979 | -34.0506 | -2.3963
1.404 | -34.0511 | -2.4076
1.4101 | -34.0514 | -2.4181
1.4162 | -34.0518 | -2.4288
1.4223 | -34.0523 | -2.4402
1.4284 | -34.0527 | -2.4504
1.4345 | -34.0529 | -2.4619
1.4406 | -34.0532 | -2.4751
1.4467 | -34.0536 | -2.4866
1.4528 | -34.054 | -2.4979
1.4589 | -34.0545 | -2.5097
1.465 | -34.0547 | -2.5206
1.4711 | -34.0552 | -2.5319
1.4772 | -34.0556 | -2.5449
1.4833 | -34.0561 | -2.5577
1.4894 | -34.0565 | -2.5697
1.4955 | -34.057 | -2.5815
1.5016 | -34.0575 | -2.5938
1.5078 | -34.0579 | -2.6073
1.5139 | -34.0583 | -2.62
1.52 | -34.0587 | -2.6324
1.5261 | -34.059 | -2.6448
1.5322 | -34.0593 | -2.6584
1.5383 | -34.0596 | -2.6707
1.5444 | -34.0601 | -2.683
1.5505 | -34.0605 | -2.6961
1.5566 | -34.061 | -2.7088
1.5627 | -34.0615 | -2.7222
1.5688 | -34.062 | -2.7354
1.5749 | -34.0623 | -2.7477
1.581 | -34.0626 | -2.7607
1.5871 | -34.063 | -2.7741
1.5932 | -34.0633 | -2.7866
1.5993 | -34.0637 | -2.8004
1.6054 | -34.064 | -2.8132
1.6115 | -34.0644 | -2.8269
1.6176 | -34.0647 | -2.8392
1.6237 | -34.0649 | -2.8524
1.6298 | -34.0651 | -2.8647
1.6359 | -34.0655 | -2.8772
1.642 | -34.0659 | -2.8889
1.6482 | -34.0661 | -2.9015
1.6543 | -34.0664 | -2.9148
1.6604 | -34.0667 | -2.9266
1.6665 | -34.0668 | -2.9388
1.6726 | -34.067 | -2.9511
1.6787 | -34.0673 | -2.9633
1.6848 | -34.0672 | -2.9756
1.6909 | -34.0675 | -2.988
1.697 | -34.0677 | -3
1.7031 | -34.0677 | -3.0123
1.7092 | -34.0677 | -3.0242
1.7153 | -34.0679 | -3.0359
1.7214 | -34.0677 | -3.048
1.7275 | -34.0677 | -3.0603
1.7336 | -34.0676 | -3.0729
1.7397 | -34.0678 | -3.085
1.7458 | -34.0678 | -3.097
1.7519 | -34.0678 | -3.1089
1.758 | -34.0678 | -3.1207
1.7641 | -34.068 | -3.1322
1.7702 | -34.0679 | -3.1448
1.7763 | -34.0679 | -3.1561
1.7824 | -34.0679 | -3.1683
1.7885 | -34.0679 | -3.1806
1.7947 | -34.068 | -3.1924
1.8008 | -34.068 | -3.2036
1.8069 | -34.068 | -3.215
1.813 | -34.0681 | -3.227
1.8191 | -34.0681 | -3.2394
1.8252 | -34.0681 | -3.2519
1.8313 | -34.0683 | -3.264
1.8374 | -34.0684 | -3.2757
1.8435 | -34.0685 | -3.2875
1.8496 | -34.0685 | -3.3003
1.8557 | -34.0687 | -3.3128
1.8618 | -34.0689 | -3.3248
1.8679 | -34.069 | -3.3367
1.874 | -34.0691 | -3.3486
1.8801 | -34.0692 | -3.3608
1.8862 | -34.0695 | -3.3726
1.8923 | -34.0696 | -3.3846
1.8984 | -34.0697 | -3.3966
1.9045 | -34.0698 | -3.4081
1.9106 | -34.0698 | -3.4202
1.9167 | -34.0699 | -3.4323
1.9228 | -34.07 | -3.444
1.9289 | -34.0701 | -3.4558
1.9351 | -34.0703 | -3.4668
1.9412 | -34.0705 | -3.478
1.9473 | -34.0707 | -3.4899
1.9534 | -34.0709 | -3.5013
1.9595 | -34.0711 | -3.5131
1.9656 | -34.0713 | -3.5253
1.9717 | -34.0715 | -3.5369
1.9778 | -34.0718 | -3.5483
1.9839 | -34.072 | -3.5597
1.99 | -34.0722 | -3.5708
1.9961 | -34.0723 | -3.5829
2.0022 | -34.0725 | -3.5943
2.0083 | -34.0726 | -3.6056
2.0144 | -34.0728 | -3.6169
2.0205 | -34.0731 | -3.6286
2.0266 | -34.0734 | -3.64
2.0327 | -34.0736 | -3.6516
2.0388 | -34.074 | -3.6631
2.0449 | -34.0742 | -3.6745
2.051 | -34.0744 | -3.6858
2.0571 | -34.0747 | -3.6971
2.0632 | -34.0749 | -3.7086
2.0693 | -34.0751 | -3.7206
2.0754 | -34.0754 | -3.7319
2.0816 | -34.0757 | -3.744
2.0877 | -34.0759 | -3.7554
2.0938 | -34.0762 | -3.7674
2.0999 | -34.0764 | -3.779
2.106 | -34.0766 | -3.7906
2.1121 | -34.0768 | -3.8024
2.1182 | -34.077 | -3.8132
2.1243 | -34.0773 | -3.8241
2.1304 | -34.0775 | -3.8352
2.1365 | -34.0777 | -3.8469
2.1426 | -34.078 | -3.8587
2.1487 | -34.0782 | -3.8699
2.1548 | -34.0786 | -3.8816
2.1609 | -34.079 | -3.893
2.167 | -34.0794 | -3.9042
2.1731 | -34.0797 | -3.916
2.1792 | -34.0799 | -3.9271
2.1853 | -34.0803 | -3.9389
2.1914 | -34.0805 | -3.9502
2.1975 | -34.0809 | -3.9617
2.2036 | -34.0811 | -3.9737
2.2097 | -34.0814 | -3.9851
2.2158 | -34.0817 | -3.9957
2.222 | -34.082 | -4.0063
2.2281 | -34.0822 | -4.0176
2.2342 | -34.0825 | -4.0289
2.2403 | -34.0827 | -4.0396
2.2464 | -34.083 | -4.0503
2.2525 | -34.0832 | -4.0616
2.2586 | -34.0836 | -4.0722
2.2647 | -34.0839 | -4.0827
2.2708 | -34.0842 | -4.0939
2.2769 | -34.0845 | -4.1046
2.283 | -34.0847 | -4.1157
2.2891 | -34.0849 | -4.1266
2.2952 | -34.0851 | -4.1371
2.3013 | -34.0853 | -4.148
2.3074 | -34.0856 | -4.1585
2.3135 | -34.0858 | -4.1698
2.3196 | -34.086 | -4.1801
2.3257 | -34.0862 | -4.1908
2.3318 | -34.0864 | -4.2019
2.3379 | -34.0865 | -4.2127
2.344 | -34.0867 | -4.224
2.3501 | -34.0869 | -4.2347
2.3562 | -34.0871 | -4.2451
2.3623 | -34.0872 | -4.2556
2.3685 | -34.0873 | -4.2661
2.3746 | -34.0875 | -4.2773
2.3807 | -34.0877 | -4.2878
2.3868 | -34.0879 | -4.2985
2.3929 | -34.0881 | -4.3092
2.399 | -34.0882 | -4.3197
2.4051 | -34.0884 | -4.3302
2.4112 | -34.0886 | -4.3409
2.4173 | -34.0888 | -4.3515
2.4234 | -34.0889 | -4.3623
2.4295 | -34.089 | -4.3731
2.4356 | -34.0892 | -4.3839
2.4417 | -34.0893 | -4.395
2.4478 | -34.0895 | -4.4058
2.4539 | -34.0897 | -4.4161
2.46 | -34.0898 | -4.4274
2.4661 | -34.0899 | -4.4385
2.4722 | -34.0901 | -4.4487
2.4783 | -34.0902 | -4.4591
2.4844 | -34.0903 | -4.4701
2.4905 | -34.0904 | -4.4811
2.4966 | -34.0905 | -4.4918
2.5027 | -34.0906 | -4.5023
2.5089 | -34.0908 | -4.5131
2.515 | -34.0908 | -4.5239
2.5211 | -34.0909 | -4.5346
2.5272 | -34.091 | -4.5454
2.5333 | -34.0911 | -4.5564
2.5394 | -34.0913 | -4.5671
2.5455 | -34.0914 | -4.5784
2.5516 | -34.0915 | -4.5887
2.5577 | -34.0917 | -4.6
2.5638 | -34.0919 | -4.6115
2.5699 | -34.0919 | -4.6228
2.576 | -34.092 | -4.6342
2.5821 | -34.0921 | -4.6454
2.5882 | -34.0922 | -4.6564
2.5943 | -34.0923 | -4.6672
2.6004 | -34.0925 | -4.678
2.6065 | -34.0926 | -4.6884
2.6126 | -34.0927 | -4.6989
2.6187 | -34.0929 | -4.7093
2.6248 | -34.093 | -4.72
2.6309 | -34.0932 | -4.7313
2.637 | -34.0934 | -4.7427
2.6431 | -34.0935 | -4.7531
2.6492 | -34.0936 | -4.7643
2.6554 | -34.0938 | -4.7751
2.6615 | -34.0939 | -4.786
2.6676 | -34.0941 | -4.7972
2.6737 | -34.0944 | -4.8083
2.6798 | -34.0946 | -4.8192
2.6859 | -34.0948 | -4.8303
2.692 | -34.0949 | -4.8411
2.6981 | -34.0951 | -4.8521
2.7042 | -34.0953 | -4.8631
2.7103 | -34.0955 | -4.8737
2.7164 | -34.0958 | -4.8846
2.7225 | -34.096 | -4.8958
2.7286 | -34.0962 | -4.9067
2.7347 | -34.0964 | -4.9171
2.7408 | -34.0966 | -4.9273
2.7469 | -34.0968 | -4.9384
2.753 | -34.0971 | -4.9489
2.7591 | -34.0974 | -4.9601
2.7652 | -34.0976 | -4.9713
2.7713 | -34.0977 | -4.982
2.7774 | -34.098 | -4.9928
2.7835 | -34.0981 | -5.0037
2.7896 | -34.0983 | -5.0142
2.7958 | -34.0986 | -5.0249
2.8019 | -34.0988 | -5.0352
2.808 | -34.0991 | -5.0459
2.8141 | -34.0993 | -5.0569
2.8202 | -34.0996 | -5.0677
2.8263 | -34.0999 | -5.0785
2.8324 | -34.1002 | -5.0896
2.8385 | -34.1006 | -5.1004
2.8446 | -34.1009 | -5.111
2.8507 | -34.1012 | -5.1218
2.8568 | -34.1015 | -5.1322
2.8629 | -34.1018 | -5.1427
2.869 | -34.1021 | -5.1537
2.8751 | -34.1024 | -5.1644
2.8812 | -34.1028 | -5.1755
2.8873 | -34.103 | -5.1863
2.8934 | -34.1033 | -5.1972
2.8995 | -34.1036 | -5.2082
2.9056 | -34.1039 | -5.2192
2.9117 | -34.1042 | -5.2301
2.9178 | -34.1045 | -5.2407
2.9239 | -34.1049 | -5.2521
2.93 | -34.1053 | -5.2634
2.9361 | -34.1055 | -5.2741
2.9423 | -34.1058 | -5.2847
2.9484 | -34.1062 | -5.2956
2.9545 | -34.1066 | -5.3065
2.9606 | -34.1069 | -5.3178
2.9667 | -34.1073 | -5.3288
2.9728 | -34.1076 | -5.3399
2.9789 | -34.108 | -5.3509
2.985 | -34.1084 | -5.3621
2.9911 | -34.1087 | -5.3732
2.9972 | -34.1092 | -5.3842
3.0033 | -34.1095 | -5.3955
3.0094 | -34.1099 | -5.4065
3.0155 | -34.1103 | -5.4173
3.0216 | -34.1106 | -5.4285
3.0277 | -34.1109 | -5.4399
3.0338 | -34.1113 | -5.4512
3.0399 | -34.1117 | -5.4623
3.046 | -34.112 | -5.4737
3.0521 | -34.1123 | -5.4847
3.0582 | -34.1125 | -5.4966
3.0643 | -34.1129 | -5.5079
3.0704 | -34.1132 | -5.5196
3.0765 | -34.1135 | -5.5305
3.0827 | -34.1137 | -5.5417
3.0888 | -34.1141 | -5.5528
3.0949 | -34.1144 | -5.564
3.101 | -34.1148 | -5.5754
3.1071 | -34.1151 | -5.5869
3.1132 | -34.1154 | -5.5985
3.1193 | -34.1158 | -5.6098
3.1254 | -34.1161 | -5.621
3.1315 | -34.1165 | -5.6328
3.1376 | -34.1167 | -5.644
3.1437 | -34.1171 | -5.6554
3.1498 | -34.1174 | -5.6673
3.1559 | -34.1177 | -5.6786
3.162 | -34.118 | -5.6905
3.1681 | -34.1182 | -5.7022
3.1742 | -34.1185 | -5.7136
3.1803 | -34.1188 | -5.7251
3.1864 | -34.1192 | -5.7363
3.1925 | -34.1195 | -5.7477
3.1986 | -34.1197 | -5.7595
3.2047 | -34.12 | -5.7714
3.2108 | -34.1202 | -5.7824
3.2169 | -34.1205 | -5.7939
3.223 | -34.1208 | -5.8055
3.2292 | -34.121 | -5.8171
3.2353 | -34.1212 | -5.828
3.2414 | -34.1215 | -5.8393
3.2475 | -34.1217 | -5.8505
3.2536 | -34.122 | -5.8619
3.2597 | -34.1222 | -5.8728
3.2658 | -34.1224 | -5.8836
3.2719 | -34.1226 | -5.8946
3.278 | -34.1229 | -5.906
3.2841 | -34.1231 | -5.9171
3.2902 | -34.1234 | -5.9284
3.2963 | -34.1236 | -5.9397
3.3024 | -34.1237 | -5.9507
3.3085 | -34.124 | -5.9621
3.3146 | -34.1242 | -5.9732
3.3207 | -34.1244 | -5.9845
3.3268 | -34.1246 | -5.9959
3.3329 | -34.1248 | -6.0066
3.339 | -34.125 | -6.0174
3.3451 | -34.1252 | -6.0285
3.3512 | -34.1254 | -6.0392
3.3573 | -34.1256 | -6.0503
3.3634 | -34.1258 | -6.0614
3.3696 | -34.126 | -6.0726
3.3757 | -34.1261 | -6.0834
3.3818 | -34.1263 | -6.0942
3.3879 | -34.1265 | -6.1052
3.394 | -34.1267 | -6.1164
3.4001 | -34.1269 | -6.1272
3.4062 | -34.127 | -6.1381
3.4123 | -34.1271 | -6.1487
3.4184 | -34.1273 | -6.1598
3.4245 | -34.1275 | -6.1706
3.4306 | -34.1276 | -6.1812
3.4367 | -34.1278 | -6.1917
3.4428 | -34.128 | -6.2021
3.4489 | -34.1281 | -6.213
3.455 | -34.1282 | -6.2238
3.4611 | -34.1284 | -6.2349
3.4672 | -34.1286 | -6.2459
3.4733 | -34.1288 | -6.2569
3.4794 | -34.1289 | -6.2676
3.4855 | -34.1291 | -6.278
3.4916 | -34.1292 | -6.2884
3.4977 | -34.1294 | -6.2996
3.5038 | -34.1295 | -6.3104
3.5099 | -34.1297 | -6.3207
3.5161 | -34.1298 | -6.331
3.5222 | -34.13 | -6.3414
3.5283 | -34.1301 | -6.3521
3.5344 | -34.1302 | -6.3629
3.5405 | -34.1304 | -6.3734
3.5466 | -34.1305 | -6.3842
3.5527 | -34.1306 | -6.3949
3.5588 | -34.1308 | -6.4055
3.5649 | -34.1308 | -6.4161
3.571 | -34.131 | -6.4267
3.5771 | -34.1311 | -6.4377
3.5832 | -34.1312 | -6.4477
3.5893 | -34.1315 | -6.4583
3.5954 | -34.1316 | -6.4691
3.6015 | -34.1318 | -6.4798
3.6076 | -34.132 | -6.4904
3.6137 | -34.1321 | -6.5005
3.6198 | -34.1324 | -6.5107
3.6259 | -34.1325 | -6.5215
3.632 | -34.1327 | -6.5321
3.6381 | -34.1329 | -6.5428
3.6442 | -34.133 | -6.5533
3.6503 | -34.1332 | -6.5635
3.6565 | -34.1334 | -6.5742
3.6626 | -34.1335 | -6.5848
3.6687 | -34.1338 | -6.5955
3.6748 | -34.134 | -6.6059
3.6809 | -34.1342 | -6.6164
3.687 | -34.1344 | -6.6273
3.6931 | -34.1346 | -6.6375
3.6992 | -34.1348 | -6.6481
3.7053 | -34.135 | -6.6586
3.7114 | -34.1352 | -6.6692
3.7175 | -34.1354 | -6.6798
3.7236 | -34.1356 | -6.6899
3.7297 | -34.1358 | -6.7006
3.7358 | -34.136 | -6.7114
3.7419 | -34.1362 | -6.7222
3.748 | -34.1364 | -6.7327
3.7541 | -34.1366 | -6.7433
3.7602 | -34.1369 | -6.7537
3.7663 | -34.1371 | -6.7642
3.7724 | -34.1373 | -6.7748
3.7785 | -34.1376 | -6.7851
3.7846 | -34.1378 | -6.7955
3.7907 | -34.1381 | -6.8064
3.7969 | -34.1383 | -6.817
3.803 | -34.1386 | -6.8272
3.8091 | -34.1388 | -6.8377
3.8152 | -34.1391 | -6.8482
3.8213 | -34.1393 | -6.8589
3.8274 | -34.1395 | -6.8692
3.8335 | -34.1399 | -6.8799
3.8396 | -34.1401 | -6.8905
3.8457 | -34.1404 | -6.9011
3.8518 | -34.1408 | -6.9116
3.8579 | -34.1411 | -6.9221
3.864 | -34.1413 | -6.9328
3.8701 | -34.1416 | -6.9437
3.8762 | -34.1419 | -6.9541
3.8823 | -34.1421 | -6.9648
3.8884 | -34.1424 | -6.9754
3.8945 | -34.1428 | -6.9862
3.9006 | -34.1431 | -6.9971
3.9067 | -34.1433 | -7.0081
3.9128 | -34.1436 | -7.0188
3.9189 | -34.1439 | -7.0292
3.925 | -34.1441 | -7.0399
3.9311 | -34.1444 | -7.0501
3.9372 | -34.1447 | -7.0609
3.9434 | -34.145 | -7.0715
3.9495 | -34.1453 | -7.0823
3.9556 | -34.1456 | -7.0927
3.9617 | -34.1459 | -7.103
3.9678 | -34.1462 | -7.1133
3.9739 | -34.1465 | -7.1239
3.98 | -34.1468 | -7.1347
3.9861 | -34.147 | -7.1452
3.9922 | -34.1473 | -7.1558
3.9983 | -34.1476 | -7.1661
4.0044 | -34.1479 | -7.1763
4.0105 | -34.1482 | -7.1869
4.0166 | -34.1485 | -7.1975
4.0227 | -34.1488 | -7.2081
4.0288 | -34.149 | -7.2191
4.0349 | -34.1492 | -7.2295
4.041 | -34.1495 | -7.24
4.0471 | -34.1497 | -7.2507
4.0532 | -34.15 | -7.2609
4.0593 | -34.1503 | -7.2717
4.0654 | -34.1506 | -7.282
4.0715 | -34.1509 | -7.293
4.0776 | -34.1512 | -7.3035
4.0838 | -34.1515 | -7.3144
4.0899 | -34.1518 | -7.3248
4.096 | -34.1521 | -7.3353
4.1021 | -34.1523 | -7.3457
4.1082 | -34.1526 | -7.356
4.1143 | -34.1528 | -7.3663
4.1204 | -34.1531 | -7.3766
4.1265 | -34.1534 | -7.3869
4.1326 | -34.1537 | -7.3971
4.1387 | -34.154 | -7.4074
4.1448 | -34.1543 | -7.418
4.1509 | -34.1545 | -7.4285
4.157 | -34.1547 | -7.4394
4.1631 | -34.155 | -7.4501
4.1692 | -34.1553 | -7.4604
4.1753 | -34.1556 | -7.4709
4.1814 | -34.1558 | -7.481
4.1875 | -34.1561 | -7.4918
4.1936 | -34.1563 | -7.5024
4.1997 | -34.1566 | -7.5125
4.2058 | -34.1568 | -7.5232
4.2119 | -34.1571 | -7.5336
4.218 | -34.1573 | -7.5438
4.2241 | -34.1575 | -7.554
4.2303 | -34.1578 | -7.5643
4.2364 | -34.1581 | -7.5747
4.2425 | -34.1582 | -7.5852
4.2486 | -34.1585 | -7.5958
4.2547 | -34.1586 | -7.606
4.2608 | -34.1588 | -7.6165
4.2669 | -34.1591 | -7.6266
4.273 | -34.1593 | -7.6372
4.2791 | -34.1595 | -7.6476
4.2852 | -34.1598 | -7.6581
4.2913 | -34.1599 | -7.6683
4.2974 | -34.1602 | -7.6789
4.3035 | -34.1604 | -7.6895
4.3096 | -34.1606 | -7.7001
4.3157 | -34.1609 | -7.7103
4.3218 | -34.1611 | -7.7206
4.3279 | -34.1613 | -7.731
4.334 | -34.1615 | -7.7413
4.3401 | -34.1617 | -7.7513
4.3462 | -34.162 | -7.7618
4.3523 | -34.1623 | -7.7718
4.3584 | -34.1624 | -7.7817
4.3645 | -34.1626 | -7.792
4.3707 | -34.1628 | -7.8021
4.3768 | -34.1631 | -7.8123
4.3829 | -34.1633 | -7.823
4.389 | -34.1635 | -7.8333
4.3951 | -34.1637 | -7.8434
4.4012 | -34.1639 | -7.854
4.4073 | -34.1641 | -7.8642
4.4134 | -34.1643 | -7.8744
4.4195 | -34.1646 | -7.8846
4.4256 | -34.1648 | -7.8947
4.4317 | -34.1651 | -7.9049
4.4378 | -34.1652 | -7.9149
4.4439 | -34.1654 | -7.9249
4.45 | -34.1656 | -7.9352
4.4561 | -34.1657 | -7.9454
4.4622 | -34.1659 | -7.9557
4.4683 | -34.1662 | -7.9659
4.4744 | -34.1664 | -7.9764
4.4805 | -34.1666 | -7.9863
4.4866 | -34.1668 | -7.9966
4.4927 | -34.167 | -8.007
4.4988 | -34.1672 | -8.0172
4.5049 | -34.1674 | -8.0271
4.511 | -34.1677 | -8.037
4.5172 | -34.1679 | -8.0472
4.5233 | -34.1681 | -8.0577
4.5294 | -34.1683 | -8.0678
4.5355 | -34.1685 | -8.078
4.5416 | -34.1687 | -8.0882
4.5477 | -34.1689 | -8.0984
4.5538 | -34.1691 | -8.1082
4.5599 | -34.1693 | -8.1183
4.566 | -34.1695 | -8.1285
4.5721 | -34.1697 | -8.1387
4.5782 | -34.17 | -8.1491
4.5843 | -34.1702 | -8.1594
4.5904 | -34.1704 | -8.1696
4.5965 | -34.1707 | -8.1797
4.6026 | -34.1709 | -8.1903
4.6087 | -34.1711 | -8.2004
4.6148 | -34.1713 | -8.2108
4.6209 | -34.1715 | -8.221
4.627 | -34.1718 | -8.2313
4.6331 | -34.1719 | -8.2414
4.6392 | -34.1721 | -8.2517
4.6453 | -34.1723 | -8.2619
4.6514 | -34.1726 | -8.2721
4.6576 | -34.1728 | -8.2826
4.6637 | -34.1729 | -8.2925
4.6698 | -34.1731 | -8.3026
4.6759 | -34.1733 | -8.3131
4.682 | -34.1735 | -8.3232
4.6881 | -34.1737 | -8.3334
4.6942 | -34.1739 | -8.3435
4.7003 | -34.1741 | -8.3538
4.7064 | -34.1743 | -8.3638
4.7125 | -34.1746 | -8.3742
4.7186 | -34.1748 | -8.3842
4.7247 | -34.175 | -8.3944
4.7308 | -34.1752 | -8.4046
4.7369 | -34.1755 | -8.4149
4.743 | -34.1757 | -8.4254
4.7491 | -34.1759 | -8.4358
4.7552 | -34.1761 | -8.4461
4.7613 | -34.1763 | -8.4564
4.7674 | -34.1765 | -8.4666
4.7735 | -34.1767 | -8.4772
4.7796 | -34.1769 | -8.4876
4.7857 | -34.1772 | -8.4979
4.7918 | -34.1774 | -8.508
4.7979 | -34.1776 | -8.5181
4.8041 | -34.1778 | -8.5284
4.8102 | -34.1781 | -8.5384
4.8163 | -34.1783 | -8.5485
4.8224 | -34.1785 | -8.5588
4.8285 | -34.1787 | -8.5688
4.8346 | -34.179 | -8.5788
4.8407 | -34.1792 | -8.5891
4.8468 | -34.1794 | -8.5996
4.8529 | -34.1797 | -8.6101
4.859 | -34.1799 | -8.6206
4.8651 | -34.1801 | -8.6306
4.8712 | -34.1803 | -8.6407
4.8773 | -34.1805 | -8.651
4.8834 | -34.1808 | -8.6612
4.8895 | -34.181 | -8.6713
4.8956 | -34.1812 | -8.6815
4.9017 | -34.1814 | -8.6916
4.9078 | -34.1816 | -8.7021
4.9139 | -34.1818 | -8.712
4.92 | -34.1821 | -8.7219
4.9261 | -34.1824 | -8.7318
4.9322 | -34.1826 | -8.7421
4.9383 | -34.1828 | -8.7523
4.9445 | -34.1831 | -8.7626
4.9506 | -34.1833 | -8.7726
4.9567 | -34.1835 | -8.7826
4.9628 | -34.1837 | -8.7928
4.9689 | -34.184 | -8.8033
4.975 | -34.1842 | -8.8138
4.9811 | -34.1844 | -8.8238
4.9872 | -34.1847 | -8.8339
4.9933 | -34.185 | -8.8444
4.9994 | -34.1853 | -8.855
5.0055 | -34.1855 | -8.8652
5.0116 | -34.1858 | -8.8758
5.0177 | -34.1861 | -8.8863
5.0238 | -34.1863 | -8.8968
5.0299 | -34.1866 | -8.9074
5.036 | -34.1869 | -8.9176
5.0421 | -34.1871 | -8.9278
5.0482 | -34.1874 | -8.9381
5.0543 | -34.1876 | -8.9481
5.0604 | -34.1879 | -8.9583
5.0665 | -34.1882 | -8.9684
5.0726 | -34.1884 | -8.9789
5.0787 | -34.1887 | -8.9894
5.0848 | -34.189 | -8.9997
5.091 | -34.1892 | -9.0098
5.0971 | -34.1895 | -9.0197
5.1032 | -34.1897 | -9.0303
5.1093 | -34.19 | -9.0406
5.1154 | -34.1902 | -9.0506
5.1215 | -34.1905 | -9.0607
5.1276 | -34.1907 | -9.0708
5.1337 | -34.1911 | -9.0808
5.1398 | -34.1913 | -9.0908
5.1459 | -34.1916 | -9.1013
5.152 | -34.1918 | -9.1117
5.1581 | -34.1921 | -9.1221
5.1642 | -34.1924 | -9.1325
5.1703 | -34.1926 | -9.1427
5.1764 | -34.1929 | -9.1529
5.1825 | -34.1932 | -9.1629
5.1886 | -34.1935 | -9.1729
5.1947 | -34.1937 | -9.1832
5.2008 | -34.194 | -9.1936
5.2069 | -34.1943 | -9.2041
5.213 | -34.1945 | -9.2141
5.2191 | -34.1948 | -9.2244
5.2252 | -34.1951 | -9.2347
5.2314 | -34.1953 | -9.2452
5.2375 | -34.1956 | -9.2552
5.2436 | -34.1959 | -9.2656
5.2497 | -34.1961 | -9.2756
5.2558 | -34.1964 | -9.2862
5.2619 | -34.1967 | -9.2967
5.268 | -34.197 | -9.3071
5.2741 | -34.1973 | -9.3175
5.2802 | -34.1976 | -9.3281
5.2863 | -34.1979 | -9.3383
5.2924 | -34.1982 | -9.3488
5.2985 | -34.1985 | -9.3592
5.3046 | -34.1988 | -9.3695
5.3107 | -34.199 | -9.38
5.3168 | -34.1993 | -9.3904
5.3229 | -34.1996 | -9.4007
5.329 | -34.1999 | -9.4111
5.3351 | -34.2002 | -9.4211
5.3412 | -34.2004 | -9.4314
5.3473 | -34.2007 | -9.4418
5.3534 | -34.201 | -9.4523
5.3595 | -34.2013 | -9.4627
5.3656 | -34.2016 | -9.473
5.3717 | -34.2019 | -9.4836
5.3779 | -34.2022 | -9.4942
5.384 | -34.2025 | -9.5046
5.3901 | -34.2028 | -9.5151
5.3962 | -34.203 | -9.5253
5.4023 | -34.2034 | -9.5358
5.4084 | -34.2037 | -9.5461
5.4145 | -34.204 | -9.5563
5.4206 | -34.2043 | -9.5667
5.4267 | -34.2046 | -9.5773
5.4328 | -34.2048 | -9.5877
5.4389 | -34.2051 | -9.5979
5.445 | -34.2054 | -9.6086
5.4511 | -34.2057 | -9.6187
5.4572 | -34.206 | -9.6291
5.4633 | -34.2063 | -9.6395
5.4694 | -34.2066 | -9.6499
5.4755 | -34.2069 | -9.66
5.4816 | -34.2072 | -9.6704
5.4877 | -34.2075 | -9.6807
5.4938 | -34.2078 | -9.6908
5.4999 | -34.2081 | -9.7011
5.506 | -34.2084 | -9.7114
5.5121 | -34.2087 | -9.7217
5.5183 | -34.209 | -9.732
5.5244 | -34.2092 | -9.7426
5.5305 | -34.2096 | -9.7528
5.5366 | -34.2098 | -9.763
5.5427 | -34.2101 | -9.7735
5.5488 | -34.2104 | -9.7839
5.5549 | -34.2107 | -9.7943
5.561 | -34.211 | -9.8049
5.5671 | -34.2113 | -9.8151
5.5732 | -34.2116 | -9.8257
5.5793 | -34.2119 | -9.8364
5.5854 | -34.2122 | -9.8467
5.5915 | -34.2125 | -9.8574
5.5976 | -34.2128 | -9.8678
5.6037 | -34.2131 | -9.8783
5.6098 | -34.2134 | -9.8886
5.6159 | -34.2137 | -9.8989
5.622 | -34.214 | -9.9094
5.6281 | -34.2143 | -9.9198
5.6342 | -34.2146 | -9.9305
5.6403 | -34.2149 | -9.9409
5.6464 | -34.2152 | -9.9514
5.6525 | -34.2155 | -9.9619
5.6586 | -34.2158 | -9.972
5.6648 | -34.2161 | -9.9825
5.6709 | -34.2165 | -9.993
5.677 | -34.2167 | -10.0034
5.6831 | -34.217 | -10.0139
5.6892 | -34.2173 | -10.0245
5.6953 | -34.2176 | -10.0349
5.7014 | -34.2179 | -10.0454
5.7075 | -34.2182 | -10.0558
5.7136 | -34.2185 | -10.0661
5.7197 | -34.2188 | -10.0765
5.7258 | -34.2191 | -10.0868
5.7319 | -34.2194 | -10.0972
5.738 | -34.2197 | -10.1076
5.7441 | -34.22 | -10.118
5.7502 | -34.2204 | -10.1281
5.7563 | -34.2207 | -10.1385
5.7624 | -34.221 | -10.1491
5.7685 | -34.2213 | -10.1595
5.7746 | -34.2216 | -10.1702
5.7807 | -34.2218 | -10.1806
5.7868 | -34.2221 | -10.1912
5.7929 | -34.2224 | -10.2015
5.799 | -34.2227 | -10.2118
5.8052 | -34.223 | -10.2221
5.8113 | -34.2233 | -10.2324
5.8174 | -34.2236 | -10.2427
5.8235 | -34.2239 | -10.253
5.8296 | -34.2242 | -10.2636
5.8357 | -34.2244 | -10.274
5.8418 | -34.2248 | -10.2845
5.8479 | -34.225 | -10.295
5.854 | -34.2253 | -10.3055
5.8601 | -34.2256 | -10.3157
5.8662 | -34.2259 | -10.3263
5.8723 | -34.2262 | -10.3365
5.8784 | -34.2265 | -10.3468
5.8845 | -34.2268 | -10.3571
5.8906 | -34.2271 | -10.3676
5.8967 | -34.2274 | -10.3778
5.9028 | -34.2277 | -10.388
5.9089 | -34.228 | -10.3983
5.915 | -34.2283 | -10.4085
5.9211 | -34.2286 | -10.4188
5.9272 | -34.2289 | -10.429
5.9333 | -34.2292 | -10.4392
5.9394 | -34.2295 | -10.4496
5.9455 | -34.2298 | -10.4598
5.9517 | -34.2302 | -10.4699
5.9578 | -34.2305 | -10.4804
5.9639 | -34.2308 | -10.4904
5.97 | -34.2311 | -10.5008
5.9761 | -34.2314 | -10.5111
5.9822 | -34.2317 | -10.5215
5.9883 | -34.232 | -10.5318
5.9944 | -34.2323 | -10.5422
6.0005 | -34.2327 | -10.5527
6.0066 | -34.233 | -10.5628
6.0127 | -34.2333 | -10.5729
6.0188 | -34.2336 | -10.5835
6.0249 | -34.2339 | -10.5939
6.031 | -34.2342 | -10.6041
6.0371 | -34.2345 | -10.6144
6.0432 | -34.2348 | -10.6246
6.0493 | -34.2351 | -10.6347
6.0554 | -34.2354 | -10.6449
6.0615 | -34.2357 | -10.6552
6.0676 | -34.2361 | -10.6657
6.0737 | -34.2364 | -10.6761
6.0798 | -34.2366 | -10.6864
6.0859 | -34.237 | -10.6966
6.0921 | -34.2373 | -10.7067
6.0982 | -34.2376 | -10.717
6.1043 | -34.2379 | -10.7271
6.1104 | -34.2383 | -10.7376
6.1165 | -34.2386 | -10.7474
6.1226 | -34.2389 | -10.7576
6.1287 | -34.2392 | -10.7679
6.1348 | -34.2395 | -10.7781
6.1409 | -34.2398 | -10.7883
6.147 | -34.2401 | -10.7985
6.1531 | -34.2404 | -10.8087
6.1592 | -34.2407 | -10.8188
6.1653 | -34.241 | -10.8289
6.1714 | -34.2413 | -10.8392
6.1775 | -34.2417 | -10.8495
6.1836 | -34.242 | -10.8598
6.1897 | -34.2423 | -10.87
6.1958 | -34.2426 | -10.8804
6.2019 | -34.2429 | -10.8908
6.208 | -34.2433 | -10.901
6.2141 | -34.2436 | -10.9111
6.2202 | -34.2439 | -10.9214
6.2263 | -34.2442 | -10.9315
6.2325 | -34.2445 | -10.9415
6.2386 | -34.2448 | -10.9514
6.2447 | -34.2451 | -10.9619
6.2508 | -34.2454 | -10.972
6.2569 | -34.2458 | -10.9822
6.263 | -34.2461 | -10.9924
6.2691 | -34.2464 | -11.0025
6.2752 | -34.2467 | -11.0127
6.2813 | -34.247 | -11.0229
6.2874 | -34.2473 | -11.033
6.2935 | -34.2476 | -11.0431
6.2996 | -34.2479 | -11.0533
6.3057 | -34.2482 | -11.0635
6.3118 | -34.2485 | -11.0737
6.3179 | -34.2488 | -11.0839
6.324 | -34.2491 | -11.094
6.3301 | -34.2494 | -11.1041
6.3362 | -34.2497 | -11.1142
6.3423 | -34.2501 | -11.1245
6.3484 | -34.2504 | -11.1347
6.3545 | -34.2507 | -11.1445
6.3606 | -34.251 | -11.1544
6.3667 | -34.2512 | -11.1645
6.3728 | -34.2516 | -11.1747
6.379 | -34.2519 | -11.1848
6.3851 | -34.2522 | -11.1947
6.3912 | -34.2525 | -11.2048
6.3973 | -34.2528 | -11.215
6.4034 | -34.2531 | -11.2249
6.4095 | -34.2534 | -11.235
6.4156 | -34.2537 | -11.2452
6.4217 | -34.254 | -11.2552
6.4278 | -34.2543 | -11.2654
6.4339 | -34.2546 | -11.2754
6.44 | -34.2549 | -11.2855
6.4461 | -34.2552 | -11.2958
6.4522 | -34.2555 | -11.3057
6.4583 | -34.2558 | -11.3159
6.4644 | -34.2561 | -11.3261
6.4705 | -34.2564 | -11.3359
6.4766 | -34.2567 | -11.346
6.4827 | -34.257 | -11.3561
6.4888 | -34.2573 | -11.366
6.4949 | -34.2576 | -11.3759
6.501 | -34.2579 | -11.386
6.5071 | -34.2582 | -11.3958
6.5132 | -34.2585 | -11.4059
6.5194 | -34.2588 | -11.4161
6.5255 | -34.2591 | -11.4259
6.5316 | -34.2594 | -11.4361
6.5377 | -34.2597 | -11.4461
6.5438 | -34.26 | -11.456
6.5499 | -34.2603 | -11.4659
6.556 | -34.2606 | -11.4759
6.5621 | -34.2609 | -11.4862
6.5682 | -34.2612 | -11.4963
6.5743 | -34.2616 | -11.5062
6.5804 | -34.2618 | -11.5163
6.5865 | -34.2622 | -11.5264
6.5926 | -34.2625 | -11.5364
6.5987 | -34.2628 | -11.5461
6.6048 | -34.2631 | -11.5562
6.6109 | -34.2634 | -11.5661
6.617 | -34.2637 | -11.5763
6.6231 | -34.264 | -11.5863
6.6292 | -34.2643 | -11.5963
6.6353 | -34.2646 | -11.6065
6.6414 | -34.2649 | -11.6165
6.6475 | -34.2652 | -11.6264
6.6536 | -34.2655 | -11.6363
6.6597 | -34.2658 | -11.6461
6.6659 | -34.2661 | -11.6561
6.672 | -34.2664 | -11.666
6.6781 | -34.2667 | -11.676
6.6842 | -34.267 | -11.6858
6.6903 | -34.2674 | -11.6958
6.6964 | -34.2677 | -11.7058
6.7025 | -34.268 | -11.7156
6.7086 | -34.2683 | -11.7256
6.7147 | -34.2686 | -11.7356
6.7208 | -34.2689 | -11.7457
6.7269 | -34.2692 | -11.7556
6.733 | -34.2695 | -11.7654
6.7391 | -34.2698 | -11.7753
6.7452 | -34.2701 | -11.7851
6.7513 | -34.2704 | -11.7948
6.7574 | -34.2707 | -11.8048
6.7635 | -34.271 | -11.8146
6.7696 | -34.2713 | -11.8244
6.7757 | -34.2717 | -11.8344
6.7818 | -34.272 | -11.8443
6.7879 | -34.2723 | -11.854
6.794 | -34.2726 | -11.8638
6.8001 | -34.2729 | -11.8737
6.8063 | -34.2732 | -11.8836
6.8124 | -34.2735 | -11.8936
6.8185 | -34.2738 | -11.9037
6.8246 | -34.2741 | -11.9134
6.8307 | -34.2744 | -11.9232
6.8368 | -34.2747 | -11.9333
6.8429 | -34.275 | -11.9432
6.849 | -34.2753 | -11.9532
6.8551 | -34.2757 | -11.9629
6.8612 | -34.276 | -11.9728
6.8673 | -34.2762 | -11.9826
6.8734 | -34.2766 | -11.9926
6.8795 | -34.2769 | -12.0025
6.8856 | -34.2772 | -12.0124
6.8917 | -34.2775 | -12.0222
6.8978 | -34.2778 | -12.0319
6.9039 | -34.2781 | -12.0415
6.91 | -34.2784 | -12.0514
6.9161 | -34.2787 | -12.0611
6.9222 | -34.279 | -12.0709
6.9283 | -34.2794 | -12.0805
6.9344 | -34.2797 | -12.0904
6.9405 | -34.28 | -12.1003
6.9466 | -34.2804 | -12.1103
6.9528 | -34.2807 | -12.1203
6.9589 | -34.281 | -12.1303
6.965 | -34.2813 | -12.1402
6.9711 | -34.2817 | -12.1501
6.9772 | -34.282 | -12.16
6.9833 | -34.2823 | -12.1698
6.9894 | -34.2826 | -12.1799
6.9955 | -34.283 | -12.1895
7.0016 | -34.2833 | -12.1993
7.0077 | -34.2836 | -12.2091
7.0138 | -34.2839 | -12.2189
7.0199 | -34.2842 | -12.2286
7.026 | -34.2845 | -12.2384
7.0321 | -34.2849 | -12.248
7.0382 | -34.2852 | -12.2576
7.0443 | -34.2855 | -12.2676
7.0504 | -34.2858 | -12.2775
7.0565 | -34.2862 | -12.2873
7.0626 | -34.2865 | -12.2969
7.0687 | -34.2868 | -12.3063
7.0748 | -34.2871 | -12.3159
7.0809 | -34.2875 | -12.3257
7.087 | -34.2878 | -12.3355
7.0932 | -34.2881 | -12.3454
7.0993 | -34.2884 | -12.3551
7.1054 | -34.2887 | -12.3648
7.1115 | -34.2891 | -12.3743
7.1176 | -34.2894 | -12.3841
7.1237 | -34.2897 | -12.3938
7.1298 | -34.29 | -12.4035
7.1359 | -34.2903 | -12.4133
7.142 | -34.2906 | -12.4228
7.1481 | -34.291 | -12.4325
7.1542 | -34.2913 | -12.4423
7.1603 | -34.2916 | -12.452
7.1664 | -34.292 | -12.4617
7.1725 | -34.2923 | -12.4716
7.1786 | -34.2926 | -12.4813
7.1847 | -34.2929 | -12.4911
7.1908 | -34.2933 | -12.501
7.1969 | -34.2935 | -12.5109
7.203 | -34.2939 | -12.5205
7.2091 | -34.2942 | -12.5302
7.2152 | -34.2946 | -12.5399
7.2213 | -34.2949 | -12.5495
7.2274 | -34.2952 | -12.5592
7.2335 | -34.2955 | -12.5689
7.2397 | -34.2959 | -12.5786
7.2458 | -34.2962 | -12.5882
7.2519 | -34.2966 | -12.5979
7.258 | -34.2969 | -12.6078
7.2641 | -34.2972 | -12.6175
7.2702 | -34.2975 | -12.6274
7.2763 | -34.2979 | -12.6372
7.2824 | -34.2982 | -12.647
7.2885 | -34.2985 | -12.6568
7.2946 | -34.2989 | -12.6666
7.3007 | -34.2992 | -12.6764
7.3068 | -34.2995 | -12.6863
7.3129 | -34.2998 | -12.6962
7.319 | -34.3002 | -12.706
7.3251 | -34.3005 | -12.7156
7.3312 | -34.3008 | -12.7253
7.3373 | -34.3012 | -12.7348
7.3434 | -34.3015 | -12.7445
7.3495 | -34.3019 | -12.7541
7.3556 | -34.3022 | -12.7639
7.3617 | -34.3025 | -12.7738
7.3678 | -34.3028 | -12.7835
7.3739 | -34.3032 | -12.7934
7.3801 | -34.3035 | -12.8034
7.3862 | -34.3038 | -12.8131
7.3923 | -34.3042 | -12.8229
7.3984 | -34.3045 | -12.8324
7.4045 | -34.3049 | -12.8421
7.4106 | -34.3052 | -12.8518
7.4167 | -34.3055 | -12.8616
7.4228 | -34.3059 | -12.8715
7.4289 | -34.3062 | -12.8812
7.435 | -34.3066 | -12.8907
7.4411 | -34.3069 | -12.9006
7.4472 | -34.3072 | -12.9104
7.4533 | -34.3076 | -12.9199
7.4594 | -34.3079 | -12.9299
7.4655 | -34.3083 | -12.9394
7.4716 | -34.3086 | -12.9493
7.4777 | -34.3089 | -12.959
7.4838 | -34.3093 | -12.9687
7.4899 | -34.3096 | -12.9784
7.496 | -34.3099 | -12.988
7.5021 | -34.3103 | -12.9977
7.5082 | -34.3106 | -13.0074
7.5143 | -34.311 | -13.0172
7.5204 | -34.3113 | -13.027
7.5266 | -34.3116 | -13.0367
7.5327 | -34.312 | -13.0463
7.5388 | -34.3123 | -13.0564
7.5449 | -34.3127 | -13.0662
7.551 | -34.313 | -13.0759
7.5571 | -34.3134 | -13.0857
7.5632 | -34.3137 | -13.0955
7.5693 | -34.3141 | -13.1052
7.5754 | -34.3144 | -13.1148
7.5815 | -34.3147 | -13.1246
7.5876 | -34.3151 | -13.1345
7.5937 | -34.3155 | -13.1444
7.5998 | -34.3158 | -13.1541
7.6059 | -34.3162 | -13.164
7.612 | -34.3165 | -13.1737
7.6181 | -34.3169 | -13.1834
7.6242 | -34.3172 | -13.1934
7.6303 | -34.3176 | -13.2034
7.6364 | -34.3179 | -13.2132
7.6425 | -34.3182 | -13.223
7.6486 | -34.3185 | -13.2328
7.6547 | -34.3189 | -13.2424
7.6608 | -34.3193 | -13.2522
7.667 | -34.3196 | -13.2621
7.6731 | -34.3199 | -13.2719
7.6792 | -34.3203 | -13.2819
7.6853 | -34.3207 | -13.2918
7.6914 | -34.321 | -13.3016
7.6975 | -34.3214 | -13.3115
7.7036 | -34.3217 | -13.3214
7.7097 | -34.3221 | -13.3313
7.7158 | -34.3224 | -13.3411
7.7219 | -34.3228 | -13.3508
7.728 | -34.3231 | -13.3607
7.7341 | -34.3235 | -13.3706
7.7402 | -34.3238 | -13.3805
7.7463 | -34.3242 | -13.3903
7.7524 | -34.3245 | -13.4004
7.7585 | -34.3249 | -13.4102
7.7646 | -34.3252 | -13.42
7.7707 | -34.3256 | -13.43
7.7768 | -34.3259 | -13.44
7.7829 | -34.3263 | -13.45
7.789 | -34.3266 | -13.4601
7.7951 | -34.3269 | -13.4699
7.8012 | -34.3273 | -13.4797
7.8073 | -34.3277 | -13.4898
7.8135 | -34.328 | -13.4996
7.8196 | -34.3284 | -13.5097
7.8257 | -34.3287 | -13.5197
7.8318 | -34.3291 | -13.5296
7.8379 | -34.3294 | -13.5394
7.844 | -34.3298 | -13.5492
7.8501 | -34.3301 | -13.5591
7.8562 | -34.3305 | -13.5692
7.8623 | -34.3309 | -13.579
7.8684 | -34.3312 | -13.5889
7.8745 | -34.3316 | -13.5988
7.8806 | -34.3319 | -13.6089
7.8867 | -34.3323 | -13.6186
7.8928 | -34.3326 | -13.6283
7.8989 | -34.333 | -13.6382
7.905 | -34.3333 | -13.6481
7.9111 | -34.3337 | -13.6581
7.9172 | -34.334 | -13.6679
7.9233 | -34.3344 | -13.6779
7.9294 | -34.3348 | -13.6878
7.9355 | -34.3351 | -13.6975
7.9416 | -34.3354 | -13.7074
7.9477 | -34.3358 | -13.7172
7.9539 | -34.3361 | -13.727
7.96 | -34.3365 | -13.7371
7.9661 | -34.3369 | -13.747
7.9722 | -34.3372 | -13.7568
7.9783 | -34.3376 | -13.7669
7.9844 | -34.338 | -13.7768
7.9905 | -34.3383 | -13.7868
7.9966 | -34.3387 | -13.7968
8.0027 | -34.339 | -13.8068
8.0088 | -34.3394 | -13.8166
8.0149 | -34.3397 | -13.8266
8.021 | -34.3401 | -13.8366
8.0271 | -34.3404 | -13.8467
8.0332 | -34.3408 | -13.8566
8.0393 | -34.3412 | -13.8666
8.0454 | -34.3415 | -13.8768
8.0515 | -34.3419 | -13.8867
8.0576 | -34.3423 | -13.8968
8.0637 | -34.3426 | -13.9067
8.0698 | -34.3429 | -13.9165
8.0759 | -34.3433 | -13.9263
8.082 | -34.3436 | -13.9361
8.0881 | -34.344 | -13.9461
8.0942 | -34.3444 | -13.9561
8.1004 | -34.3448 | -13.9661
8.1065 | -34.3451 | -13.9759
8.1126 | -34.3454 | -13.9858
8.1187 | -34.3458 | -13.9959
8.1248 | -34.3461 | -14.0056
8.1309 | -34.3465 | -14.0154
8.137 | -34.3469 | -14.0252
8.1431 | -34.3472 | -14.0353
8.1492 | -34.3476 | -14.0452
8.1553 | -34.3479 | -14.0551
8.1614 | -34.3483 | -14.0652
8.1675 | -34.3486 | -14.0751
8.1736 | -34.349 | -14.0852
8.1797 | -34.3493 | -14.0952
8.1858 | -34.3497 | -14.1052
8.1919 | -34.3501 | -14.1151
8.198 | -34.3504 | -14.1248
8.2041 | -34.3508 | -14.1348
8.2102 | -34.3511 | -14.1447
8.2163 | -34.3515 | -14.1545
8.2224 | -34.3519 | -14.1644
8.2285 | -34.3522 | -14.1741
8.2346 | -34.3526 | -14.184
8.2408 | -34.3529 | -14.1937
8.2469 | -34.3533 | -14.2037
8.253 | -34.3537 | -14.2135
8.2591 | -34.354 | -14.2234
8.2652 | -34.3544 | -14.2334
8.2713 | -34.3548 | -14.2433
8.2774 | -34.3552 | -14.2533
8.2835 | -34.3555 | -14.2634
8.2896 | -34.3559 | -14.2732
8.2957 | -34.3563 | -14.2832
8.3018 | -34.3566 | -14.2931
8.3079 | -34.357 | -14.3028
8.314 | -34.3574 | -14.3128
8.3201 | -34.3577 | -14.3226
8.3262 | -34.3581 | -14.3325
8.3323 | -34.3584 | -14.3424
8.3384 | -34.3588 | -14.3525
8.3445 | -34.3592 | -14.3623
8.3506 | -34.3596 | -14.3722
8.3567 | -34.3599 | -14.3822
8.3628 | -34.3603 | -14.3921
8.3689 | -34.3606 | -14.4019
8.375 | -34.361 | -14.4119
8.3812 | -34.3614 | -14.4217
8.3873 | -34.3617 | -14.4316
8.3934 | -34.3621 | -14.4416
8.3995 | -34.3624 | -14.4513
8.4056 | -34.3628 | -14.4612
8.4117 | -34.3632 | -14.4712
8.4178 | -34.3635 | -14.4811
8.4239 | -34.3639 | -14.4908
8.43 | -34.3642 | -14.5007
8.4361 | -34.3646 | -14.5107
8.4422 | -34.365 | -14.5205
8.4483 | -34.3653 | -14.5303
8.4544 | -34.3657 | -14.5402
8.4605 | -34.3661 | -14.55
8.4666 | -34.3664 | -14.56
8.4727 | -34.3668 | -14.5699
8.4788 | -34.3671 | -14.5798
8.4849 | -34.3675 | -14.5896
8.491 | -34.3679 | -14.5995
8.4971 | -34.3682 | -14.6092
8.5032 | -34.3686 | -14.6189
8.5093 | -34.3689 | -14.6286
8.5154 | -34.3693 | -14.6385
8.5215 | -34.3697 | -14.6483
8.5277 | -34.37 | -14.6582
8.5338 | -34.3704 | -14.6679
8.5399 | -34.3707 | -14.6778
8.546 | -34.3711 | -14.6874
8.5521 | -34.3715 | -14.6973
8.5582 | -34.3718 | -14.7071
8.5643 | -34.3722 | -14.7167
8.5704 | -34.3726 | -14.7263
8.5765 | -34.373 | -14.7362
8.5826 | -34.3733 | -14.7457
8.5887 | -34.3737 | -14.7555
8.5948 | -34.3741 | -14.7653
8.6009 | -34.3744 | -14.775
8.607 | -34.3748 | -14.7849
8.6131 | -34.3752 | -14.7948
8.6192 | -34.3756 | -14.8044
8.6253 | -34.3759 | -14.8141
8.6314 | -34.3763 | -14.8238
8.6375 | -34.3766 | -14.8337
8.6436 | -34.377 | -14.8435
8.6497 | -34.3774 | -14.8531
8.6558 | -34.3777 | -14.8628
8.6619 | -34.378 | -14.8727
8.6681 | -34.3784 | -14.8826
8.6742 | -34.3788 | -14.8923
8.6803 | -34.3791 | -14.9021
8.6864 | -34.3795 | -14.9115
8.6925 | -34.3799 | -14.9212
8.6986 | -34.3802 | -14.931
8.7047 | -34.3806 | -14.941
8.7108 | -34.3809 | -14.9507
8.7169 | -34.3813 | -14.9605
8.723 | -34.3817 | -14.97
8.7291 | -34.3821 | -14.9796
8.7352 | -34.3824 | -14.9895
8.7413 | -34.3828 | -14.9993
8.7474 | -34.3831 | -15.0089
8.7535 | -34.3835 | -15.0186
8.7596 | -34.3839 | -15.0283
8.7657 | -34.3843 | -15.0381
8.7718 | -34.3846 | -15.0479
8.7779 | -34.385 | -15.0577
8.784 | -34.3854 | -15.0675
8.7901 | -34.3857 | -15.0772
8.7962 | -34.3861 | -15.0867
8.8023 | -34.3865 | -15.0964
8.8084 | -34.3868 | -15.1061
8.8146 | -34.3872 | -15.1157
8.8207 | -34.3875 | -15.1254
8.8268 | -34.3879 | -15.1347
8.8329 | -34.3883 | -15.1444
8.839 | -34.3887 | -15.1541
8.8451 | -34.389 | -15.1637
8.8512 | -34.3894 | -15.1733
8.8573 | -34.3898 | -15.1831
8.8634 | -34.3901 | -15.1925
8.8695 | -34.3905 | -15.2024
8.8756 | -34.3908 | -15.2121
8.8817 | -34.3912 | -15.2216
8.8878 | -34.3916 | -15.2314
8.8939 | -34.392 | -15.2411
8.9 | -34.3924 | -15.2507
8.9061 | -34.3927 | -15.2603
8.9122 | -34.3931 | -15.27
8.9183 | -34.3935 | -15.2797
8.9244 | -34.3938 | -15.2894
8.9305 | -34.3942 | -15.2988
8.9366 | -34.3946 | -15.3084
8.9427 | -34.395 | -15.3181
8.9488 | -34.3954 | -15.3277
8.955 | -34.3957 | -15.3373
8.9611 | -34.3961 | -15.347
8.9672 | -34.3965 | -15.3565
8.9733 | -34.3969 | -15.366
8.9794 | -34.3973 | -15.3755
8.9855 | -34.3976 | -15.3852
8.9916 | -34.398 | -15.395
8.9977 | -34.3984 | -15.4046
9.0038 | -34.3988 | -15.414
9.0099 | -34.3992 | -15.4237
9.016 | -34.3995 | -15.4333
9.0221 | -34.3999 | -15.4428
9.0282 | -34.4003 | -15.4524
9.0343 | -34.4007 | -15.462
9.0404 | -34.4011 | -15.4716
9.0465 | -34.4015 | -15.4813
9.0526 | -34.4018 | -15.4908
9.0587 | -34.4023 | -15.5004
9.0648 | -34.4026 | -15.5099
9.0709 | -34.403 | -15.5195
9.077 | -34.4034 | -15.5291
9.0831 | -34.4038 | -15.5388
9.0892 | -34.4042 | -15.5485
9.0953 | -34.4046 | -15.5581
9.1015 | -34.405 | -15.5677
9.1076 | -34.4054 | -15.577
9.1137 | -34.4057 | -15.5866
9.1198 | -34.4061 | -15.5962
9.1259 | -34.4065 | -15.6058
9.132 | -34.4069 | -15.6155
9.1381 | -34.4073 | -15.625
9.1442 | -34.4077 | -15.6343
9.1503 | -34.4081 | -15.6439
9.1564 | -34.4085 | -15.6535
9.1625 | -34.4089 | -15.663
9.1686 | -34.4093 | -15.6725
9.1747 | -34.4097 | -15.682
9.1808 | -34.4101 | -15.6915
9.1869 | -34.4105 | -15.7009
9.193 | -34.4109 | -15.7105
9.1991 | -34.4112 | -15.7201
9.2052 | -34.4116 | -15.7296
9.2113 | -34.412 | -15.7392
9.2174 | -34.4125 | -15.7486
9.2235 | -34.4129 | -15.7582
9.2296 | -34.4133 | -15.7677
9.2357 | -34.4137 | -15.7774
9.2419 | -34.4141 | -15.7869
9.248 | -34.4145 | -15.7964
9.2541 | -34.4149 | -15.8059
9.2602 | -34.4153 | -15.8157
9.2663 | -34.4157 | -15.8249
9.2724 | -34.4161 | -15.8344
9.2785 | -34.4165 | -15.8439
9.2846 | -34.4169 | -15.8532
9.2907 | -34.4173 | -15.8627
9.2968 | -34.4177 | -15.8723
9.3029 | -34.4181 | -15.882
9.309 | -34.4185 | -15.8916
9.3151 | -34.4189 | -15.901
9.3212 | -34.4193 | -15.9103
9.3273 | -34.4197 | -15.9198
9.3334 | -34.4201 | -15.9291
9.3395 | -34.4205 | -15.9385
9.3456 | -34.4209 | -15.9479
9.3517 | -34.4213 | -15.9574
9.3578 | -34.4217 | -15.9668
9.3639 | -34.4221 | -15.9762
9.37 | -34.4225 | -15.9857
9.3761 | -34.4229 | -15.9953
9.3822 | -34.4233 | -16.0049
9.3884 | -34.4237 | -16.0141
9.3945 | -34.4241 | -16.0236
9.4006 | -34.4245 | -16.0331
9.4067 | -34.4249 | -16.0427
9.4128 | -34.4253 | -16.0523
9.4189 | -34.4257 | -16.0619
9.425 | -34.4262 | -16.0711
9.4311 | -34.4266 | -16.0805
9.4372 | -34.427 | -16.09
9.4433 | -34.4274 | -16.0994
9.4494 | -34.4278 | -16.1087
9.4555 | -34.4282 | -16.1183
9.4616 | -34.4286 | -16.1277
9.4677 | -34.429 | -16.1372
9.4738 | -34.4294 | -16.1467
9.4799 | -34.4298 | -16.1561
9.486 | -34.4302 | -16.1655
9.4921 | -34.4306 | -16.1749
9.4982 | -34.4311 | -16.1841
9.5043 | -34.4315 | -16.1936
9.5104 | -34.4319 | -16.2029
9.5165 | -34.4323 | -16.2123
9.5226 | -34.4327 | -16.2218
9.5288 | -34.4331 | -16.231
9.5349 | -34.4335 | -16.2405
9.541 | -34.4339 | -16.2498
9.5471 | -34.4343 | -16.2592
9.5532 | -34.4347 | -16.2687
9.5593 | -34.4351 | -16.2782
9.5654 | -34.4355 | -16.2875
9.5715 | -34.436 | -16.2969
9.5776 | -34.4363 | -16.3065
9.5837 | -34.4367 | -16.3159
9.5898 | -34.4371 | -16.3255
9.5959 | -34.4375 | -16.3349
9.602 | -34.4379 | -16.3442
9.6081 | -34.4383 | -16.3536
9.6142 | -34.4388 | -16.363
9.6203 | -34.4391 | -16.3724
9.6264 | -34.4395 | -16.3817
9.6325 | -34.4399 | -16.3912
9.6386 | -34.4404 | -16.4005
9.6447 | -34.4408 | -16.4099
9.6508 | -34.4412 | -16.4193
9.6569 | -34.4416 | -16.4286
9.663 | -34.442 | -16.438
9.6691 | -34.4424 | -16.4476
9.6753 | -34.4428 | -16.457
9.6814 | -34.4432 | -16.4662
9.6875 | -34.4437 | -16.4754
9.6936 | -34.4441 | -16.4848
9.6997 | -34.4445 | -16.4943
9.7058 | -34.445 | -16.5036
9.7119 | -34.4454 | -16.5131
9.718 | -34.4458 | -16.5227
9.7241 | -34.4462 | -16.5319
9.7302 | -34.4466 | -16.5414
9.7363 | -34.447 | -16.5507
9.7424 | -34.4474 | -16.5599
9.7485 | -34.4478 | -16.5693
9.7546 | -34.4482 | -16.5786
9.7607 | -34.4486 | -16.5882
9.7668 | -34.4491 | -16.5976
9.7729 | -34.4495 | -16.607
9.779 | -34.4499 | -16.6162
9.7851 | -34.4504 | -16.6257
9.7912 | -34.4508 | -16.635
9.7973 | -34.4512 | -16.6445
9.8034 | -34.4516 | -16.6537
9.8095 | -34.4521 | -16.6633
9.8157 | -34.4525 | -16.6727
9.8218 | -34.4529 | -16.6822
9.8279 | -34.4533 | -16.6914
9.834 | -34.4537 | -16.7009
9.8401 | -34.4541 | -16.7103
9.8462 | -34.4545 | -16.7194
9.8523 | -34.4549 | -16.7289
9.8584 | -34.4553 | -16.7382
9.8645 | -34.4557 | -16.7476
9.8706 | -34.4561 | -16.757
9.8767 | -34.4565 | -16.7663
9.8828 | -34.4569 | -16.7757
9.8889 | -34.4574 | -16.7849
9.895 | -34.4578 | -16.7943
9.9011 | -34.4582 | -16.8037
9.9072 | -34.4586 | -16.813
9.9133 | -34.459 | -16.8223
9.9194 | -34.4594 | -16.8316
9.9255 | -34.4598 | -16.8409
9.9316 | -34.4603 | -16.8501
9.9377 | -34.4607 | -16.8594
9.9438 | -34.4611 | -16.8689
9.9499 | -34.4615 | -16.8785
9.956 | -34.4619 | -16.8877
9.9622 | -34.4624 | -16.8971
9.9683 | -34.4628 | -16.9065
9.9744 | -34.4632 | -16.9159
9.9805 | -34.4636 | -16.9252
9.9866 | -34.464 | -16.9344
9.9927 | -34.4644 | -16.9436
9.9988 | -34.4648 | -16.9531
10.0049 | -34.4653 | -16.9624
10.011 | -34.4657 | -16.9719
10.0171 | -34.4661 | -16.9812
10.0232 | -34.4665 | -16.9905
10.0293 | -34.467 | -16.9999
10.0354 | -34.4674 | -17.0092
10.0415 | -34.4678 | -17.0185
10.0476 | -34.4682 | -17.028
10.0537 | -34.4686 | -17.0374
10.0598 | -34.4691 | -17.0467
10.0659 | -34.4694 | -17.0562
10.072 | -34.4699 | -17.0654
10.0781 | -34.4703 | -17.0747
10.0842 | -34.4707 | -17.084
10.0903 | -34.4711 | -17.0933
10.0964 | -34.4715 | -17.1024
10.1026 | -34.4719 | -17.1118
10.1087 | -34.4724 | -17.1211
10.1148 | -34.4727 | -17.1306
10.1209 | -34.4732 | -17.1398
10.127 | -34.4736 | -17.1489
10.1331 | -34.474 | -17.1582
10.1392 | -34.4744 | -17.1673
10.1453 | -34.4748 | -17.1766
10.1514 | -34.4752 | -17.1859
10.1575 | -34.4756 | -17.1954
10.1636 | -34.4761 | -17.2044
10.1697 | -34.4765 | -17.2139
10.1758 | -34.4769 | -17.223
10.1819 | -34.4773 | -17.2323
10.188 | -34.4777 | -17.2417
10.1941 | -34.4781 | -17.251
10.2002 | -34.4785 | -17.2604
10.2063 | -34.4789 | -17.2697
10.2124 | -34.4793 | -17.279
10.2185 | -34.4798 | -17.2884
10.2246 | -34.4802 | -17.2977
10.2307 | -34.4806 | -17.3071
10.2368 | -34.481 | -17.3161
10.2429 | -34.4814 | -17.3254
10.2491 | -34.4818 | -17.3348
10.2552 | -34.4822 | -17.344
10.2613 | -34.4827 | -17.3534
10.2674 | -34.4831 | -17.3627
10.2735 | -34.4835 | -17.3722
10.2796 | -34.4839 | -17.3815
10.2857 | -34.4843 | -17.3908
10.2918 | -34.4848 | -17.4002
10.2979 | -34.4852 | -17.4096
10.304 | -34.4856 | -17.4186
10.3101 | -34.486 | -17.4279
10.3162 | -34.4865 | -17.4373
10.3223 | -34.4869 | -17.4464
10.3284 | -34.4873 | -17.4557
10.3345 | -34.4877 | -17.4649
10.3406 | -34.4881 | -17.4741
10.3467 | -34.4885 | -17.4834
10.3528 | -34.4889 | -17.4926
10.3589 | -34.4893 | -17.5018
10.365 | -34.4897 | -17.5112
10.3711 | -34.4902 | -17.5206
10.3772 | -34.4906 | -17.5299
10.3833 | -34.491 | -17.5391
10.3895 | -34.4914 | -17.5483
10.3956 | -34.4919 | -17.5577
10.4017 | -34.4923 | -17.5668
10.4078 | -34.4927 | -17.5761
10.4139 | -34.4931 | -17.5855
10.42 | -34.4935 | -17.5949
10.4261 | -34.4939 | -17.6041
10.4322 | -34.4943 | -17.6133
10.4383 | -34.4947 | -17.6226
10.4444 | -34.4951 | -17.6317
10.4505 | -34.4955 | -17.641
10.4566 | -34.4959 | -17.6502
10.4627 | -34.4964 | -17.6593
10.4688 | -34.4968 | -17.6684
10.4749 | -34.4972 | -17.6777
10.481 | -34.4976 | -17.6871
10.4871 | -34.498 | -17.6963
10.4932 | -34.4984 | -17.7054
10.4993 | -34.4989 | -17.7147
10.5054 | -34.4993 | -17.724
10.5115 | -34.4997 | -17.7332
10.5176 | -34.5001 | -17.7424
10.5237 | -34.5005 | -17.7517
10.5298 | -34.5009 | -17.7609
10.536 | -34.5014 | -17.7702
10.5421 | -34.5018 | -17.7794
10.5482 | -34.5022 | -17.7886
10.5543 | -34.5026 | -17.7978
10.5604 | -34.503 | -17.807
10.5665 | -34.5034 | -17.8161
10.5726 | -34.5038 | -17.8253
10.5787 | -34.5042 | -17.8345
10.5848 | -34.5046 | -17.8434
10.5909 | -34.5051 | -17.8528
10.597 | -34.5055 | -17.8621
10.6031 | -34.5059 | -17.8716
10.6092 | -34.5063 | -17.8808
10.6153 | -34.5067 | -17.8902
10.6214 | -34.5071 | -17.8993
10.6275 | -34.5075 | -17.9086
10.6336 | -34.5079 | -17.9178
10.6397 | -34.5083 | -17.927
10.6458 | -34.5088 | -17.9362
10.6519 | -34.5092 | -17.9455
10.658 | -34.5096 | -17.9548
10.6641 | -34.51 | -17.9642
10.6702 | -34.5104 | -17.9735
10.6764 | -34.5109 | -17.9829
10.6825 | -34.5113 | -17.9919
10.6886 | -34.5117 | -18.0009
10.6947 | -34.5121 | -18.0101
10.7008 | -34.5125 | -18.0192
10.7069 | -34.5129 | -18.0283
10.713 | -34.5133 | -18.0376
10.7191 | -34.5137 | -18.0468
10.7252 | -34.5141 | -18.0557
10.7313 | -34.5145 | -18.065
10.7374 | -34.5149 | -18.0742
10.7435 | -34.5153 | -18.0836
10.7496 | -34.5157 | -18.0926
10.7557 | -34.5161 | -18.1019
10.7618 | -34.5166 | -18.1108
10.7679 | -34.517 | -18.1201
10.774 | -34.5174 | -18.1293
10.7801 | -34.5178 | -18.1387
10.7862 | -34.5182 | -18.1477
10.7923 | -34.5186 | -18.1569
10.7984 | -34.5191 | -18.1661
10.8045 | -34.5195 | -18.1752
10.8106 | -34.5199 | -18.1846
10.8168 | -34.5203 | -18.1938
10.8229 | -34.5207 | -18.203
10.829 | -34.5211 | -18.2121
10.8351 | -34.5215 | -18.221
10.8412 | -34.5219 | -18.2301
10.8473 | -34.5223 | -18.2391
10.8534 | -34.5228 | -18.2483
10.8595 | -34.5232 | -18.2575
10.8656 | -34.5236 | -18.2665
10.8717 | -34.524 | -18.2756
10.8778 | -34.5245 | -18.2848
10.8839 | -34.5249 | -18.2939
10.89 | -34.5253 | -18.3031
10.8961 | -34.5257 | -18.3122
10.9022 | -34.5261 | -18.3212
10.9083 | -34.5265 | -18.3303
10.9144 | -34.5269 | -18.3394
10.9205 | -34.5273 | -18.3485
10.9266 | -34.5277 | -18.3575
10.9327 | -34.5282 | -18.3667
10.9388 | -34.5286 | -18.3756
10.9449 | -34.529 | -18.3846
10.951 | -34.5294 | -18.3937
10.9571 | -34.5298 | -18.4029
10.9633 | -34.5302 | -18.4121
10.9694 | -34.5306 | -18.4214
10.9755 | -34.5311 | -18.4305
10.9816 | -34.5315 | -18.4397
10.9877 | -34.5319 | -18.4487
10.9938 | -34.5323 | -18.458
10.9999 | -34.5328 | -18.467
11.006 | -34.5332 | -18.4761
11.0121 | -34.5336 | -18.4852
11.0182 | -34.534 | -18.4943
11.0243 | -34.5344 | -18.5033
11.0304 | -34.5348 | -18.5125
11.0365 | -34.5352 | -18.5215
11.0426 | -34.5357 | -18.5304
11.0487 | -34.5361 | -18.5397
11.0548 | -34.5365 | -18.5487
11.0609 | -34.5369 | -18.5579
11.067 | -34.5373 | -18.567
11.0731 | -34.5378 | -18.576
11.0792 | -34.5382 | -18.5851
11.0853 | -34.5386 | -18.5943
11.0914 | -34.539 | -18.6035
11.0975 | -34.5394 | -18.6126
11.1037 | -34.5398 | -18.6216
11.1098 | -34.5402 | -18.6308
11.1159 | -34.5407 | -18.6398
11.122 | -34.5411 | -18.6488
11.1281 | -34.5415 | -18.6578
11.1342 | -34.5419 | -18.667
11.1403 | -34.5423 | -18.6761
11.1464 | -34.5427 | -18.6852
11.1525 | -34.5431 | -18.6941
11.1586 | -34.5435 | -18.7032
11.1647 | -34.5439 | -18.7122
11.1708 | -34.5443 | -18.7213
11.1769 | -34.5448 | -18.7305
11.183 | -34.5452 | -18.7395
11.1891 | -34.5456 | -18.7486
11.1952 | -34.546 | -18.7577
11.2013 | -34.5464 | -18.7669
11.2074 | -34.5468 | -18.7759
11.2135 | -34.5472 | -18.7851
11.2196 | -34.5477 | -18.7943
11.2257 | -34.5481 | -18.8034
11.2318 | -34.5485 | -18.8123
11.2379 | -34.549 | -18.8215
11.244 | -34.5494 | -18.8307
11.2502 | -34.5498 | -18.8396
11.2563 | -34.5503 | -18.8489
11.2624 | -34.5507 | -18.8579
11.2685 | -34.5511 | -18.8669
11.2746 | -34.5515 | -18.876
11.2807 | -34.552 | -18.8851
11.2868 | -34.5524 | -18.8942
11.2929 | -34.5528 | -18.9029
11.299 | -34.5532 | -18.912
11.3051 | -34.5536 | -18.9211
11.3112 | -34.5541 | -18.9302
11.3173 | -34.5545 | -18.939
11.3234 | -34.5549 | -18.948
11.3295 | -34.5553 | -18.957
11.3356 | -34.5557 | -18.9661
11.3417 | -34.5562 | -18.9751
11.3478 | -34.5566 | -18.9843
11.3539 | -34.557 | -18.9934
11.36 | -34.5574 | -19.0024
11.3661 | -34.5578 | -19.0112
11.3722 | -34.5583 | -19.02
11.3783 | -34.5587 | -19.0291
11.3844 | -34.5591 | -19.0382
11.3906 | -34.5595 | -19.0473
11.3967 | -34.56 | -19.056
11.4028 | -34.5604 | -19.0652
11.4089 | -34.5608 | -19.0741
11.415 | -34.5612 | -19.083
11.4211 | -34.5616 | -19.0921
11.4272 | -34.562 | -19.1011
11.4333 | -34.5624 | -19.1099
11.4394 | -34.5629 | -19.1187
11.4455 | -34.5633 | -19.1277
11.4516 | -34.5637 | -19.1368
11.4577 | -34.5641 | -19.1459
11.4638 | -34.5645 | -19.1548
11.4699 | -34.565 | -19.1638
11.476 | -34.5654 | -19.173
11.4821 | -34.5658 | -19.1819
11.4882 | -34.5663 | -19.1911
11.4943 | -34.5667 | -19.2002
11.5004 | -34.5671 | -19.209
11.5065 | -34.5676 | -19.2179
11.5126 | -34.568 | -19.2268
11.5187 | -34.5684 | -19.236
11.5248 | -34.5688 | -19.2451
11.5309 | -34.5692 | -19.254
11.5371 | -34.5696 | -19.2627
11.5432 | -34.5701 | -19.2718
11.5493 | -34.5705 | -19.2809
11.5554 | -34.5709 | -19.2899
11.5615 | -34.5713 | -19.2989
11.5676 | -34.5718 | -19.3078
11.5737 | -34.5722 | -19.3166
11.5798 | -34.5726 | -19.3254
11.5859 | -34.5731 | -19.3343
11.592 | -34.5735 | -19.3433
11.5981 | -34.5739 | -19.3522
11.6042 | -34.5744 | -19.3612
11.6103 | -34.5747 | -19.3702
11.6164 | -34.5752 | -19.379
11.6225 | -34.5756 | -19.3878
11.6286 | -34.576 | -19.3967
11.6347 | -34.5764 | -19.4057
11.6408 | -34.5769 | -19.4145
11.6469 | -34.5773 | -19.4234
11.653 | -34.5777 | -19.4322
11.6591 | -34.5782 | -19.4412
11.6652 | -34.5786 | -19.4501
11.6713 | -34.579 | -19.4589
11.6775 | -34.5794 | -19.4677
11.6836 | -34.5799 | -19.4766
11.6897 | -34.5803 | -19.4857
11.6958 | -34.5807 | -19.4947
11.7019 | -34.5812 | -19.5037
11.708 | -34.5816 | -19.5126
11.7141 | -34.5821 | -19.5213
11.7202 | -34.5825 | -19.5302
11.7263 | -34.5829 | -19.5391
11.7324 | -34.5833 | -19.5483
11.7385 | -34.5837 | -19.5573
11.7446 | -34.5841 | -19.5663
11.7507 | -34.5846 | -19.5749
11.7568 | -34.585 | -19.5838
11.7629 | -34.5854 | -19.5928
11.769 | -34.5858 | -19.6017
11.7751 | -34.5863 | -19.6106
11.7812 | -34.5867 | -19.6193
11.7873 | -34.5871 | -19.6282
11.7934 | -34.5876 | -19.6371
11.7995 | -34.588 | -19.646
11.8056 | -34.5884 | -19.6548
11.8117 | -34.5888 | -19.6637
11.8178 | -34.5893 | -19.6724
11.824 | -34.5897 | -19.6812
11.8301 | -34.5902 | -19.6901
11.8362 | -34.5906 | -19.6989
11.8423 | -34.591 | -19.7077
11.8484 | -34.5914 | -19.7166
11.8545 | -34.5919 | -19.7252
11.8606 | -34.5923 | -19.7339
11.8667 | -34.5928 | -19.7428
11.8728 | -34.5932 | -19.7515
11.8789 | -34.5936 | -19.7603
11.885 | -34.594 | -19.7693
11.8911 | -34.5944 | -19.778
11.8972 | -34.5949 | -19.7868
11.9033 | -34.5953 | -19.7957
11.9094 | -34.5957 | -19.8045
11.9155 | -34.5961 | -19.8133
11.9216 | -34.5966 | -19.822
11.9277 | -34.597 | -19.8308
11.9338 | -34.5974 | -19.8398
11.9399 | -34.5979 | -19.8487
11.946 | -34.5983 | -19.8575
11.9521 | -34.5987 | -19.8662
11.9582 | -34.5992 | -19.875
11.9644 | -34.5996 | -19.8837
11.9705 | -34.6 | -19.8925
11.9766 | -34.6005 | -19.9013
11.9827 | -34.6009 | -19.9101
11.9888 | -34.6014 | -19.9188
11.9949 | -34.6018 | -19.9273
12.001 | -34.6023 | -19.9361
12.0071 | -34.6027 | -19.9449
12.0132 | -34.6031 | -19.9535
12.0193 | -34.6035 | -19.9623
12.0254 | -34.6039 | -19.9711
12.0315 | -34.6044 | -19.9797
12.0376 | -34.6048 | -19.9884
12.0437 | -34.6052 | -19.9972
12.0498 | -34.6056 | -20.0059
12.0559 | -34.6061 | -20.0146
12.062 | -34.6065 | -20.0232
12.0681 | -34.607 | -20.032
12.0742 | -34.6074 | -20.0407
12.0803 | -34.6078 | -20.0496
12.0864 | -34.6083 | -20.0584
12.0925 | -34.6087 | -20.0672
12.0986 | -34.6092 | -20.0759
12.1047 | -34.6096 | -20.0847
12.1109 | -34.6101 | -20.0933
12.117 | -34.6105 | -20.1021
12.1231 | -34.611 | -20.1108
12.1292 | -34.6114 | -20.1194
12.1353 | -34.6118 | -20.1281
12.1414 | -34.6122 | -20.1369
12.1475 | -34.6127 | -20.1456
12.1536 | -34.6131 | -20.1545
12.1597 | -34.6135 | -20.1631
12.1658 | -34.614 | -20.1718
12.1719 | -34.6144 | -20.1803
12.178 | -34.6149 | -20.1892
12.1841 | -34.6153 | -20.1979
12.1902 | -34.6158 | -20.2067
12.1963 | -34.6162 | -20.2154
12.2024 | -34.6167 | -20.2239
12.2085 | -34.6171 | -20.2327
12.2146 | -34.6175 | -20.2414
12.2207 | -34.618 | -20.2501
12.2268 | -34.6184 | -20.2588
12.2329 | -34.6188 | -20.2675
12.239 | -34.6193 | -20.2759
12.2451 | -34.6198 | -20.2846
12.2513 | -34.6202 | -20.2934
12.2574 | -34.6206 | -20.3022
12.2635 | -34.6211 | -20.3109
12.2696 | -34.6215 | -20.3195
12.2757 | -34.622 | -20.328
12.2818 | -34.6224 | -20.3368
12.2879 | -34.6229 | -20.3454
12.294 | -34.6233 | -20.354
12.3001 | -34.6238 | -20.3626
12.3062 | -34.6242 | -20.3713
12.3123 | -34.6247 | -20.3797
12.3184 | -34.6251 | -20.3883
12.3245 | -34.6256 | -20.3968
12.3306 | -34.626 | -20.4055
12.3367 | -34.6264 | -20.4141
12.3428 | -34.6269 | -20.4228
12.3489 | -34.6273 | -20.4312
12.355 | -34.6278 | -20.4399
12.3611 | -34.6282 | -20.4487
12.3672 | -34.6286 | -20.4572
12.3733 | -34.6291 | -20.4659
12.3794 | -34.6296 | -20.4743
12.3855 | -34.63 | -20.4829
12.3916 | -34.6304 | -20.4915
12.3978 | -34.6309 | -20.5002
12.4039 | -34.6313 | -20.5089
12.41 | -34.6318 | -20.5175
12.4161 | -34.6322 | -20.5259
12.4222 | -34.6327 | -20.5345
12.4283 | -34.6331 | -20.5431
12.4344 | -34.6336 | -20.5518
12.4405 | -34.634 | -20.5604
12.4466 | -34.6344 | -20.569
12.4527 | -34.6349 | -20.5774
12.4588 | -34.6353 | -20.5861
12.4649 | -34.6358 | -20.5947
12.471 | -34.6362 | -20.6033
12.4771 | -34.6367 | -20.6119
12.4832 | -34.6371 | -20.6204
12.4893 | -34.6376 | -20.6289
12.4954 | -34.638 | -20.6374
12.5015 | -34.6385 | -20.646
12.5076 | -34.6389 | -20.6547
12.5137 | -34.6394 | -20.6633
12.5198 | -34.6398 | -20.6719
12.5259 | -34.6403 | -20.6805
12.532 | -34.6407 | -20.6892
12.5382 | -34.6412 | -20.6979
12.5443 | -34.6416 | -20.7064
12.5504 | -34.6421 | -20.7151
12.5565 | -34.6425 | -20.7235
12.5626 | -34.643 | -20.7323
12.5687 | -34.6434 | -20.741
12.5748 | -34.6439 | -20.7496
12.5809 | -34.6443 | -20.7582
12.587 | -34.6448 | -20.7669
12.5931 | -34.6453 | -20.7753
12.5992 | -34.6458 | -20.7839
12.6053 | -34.6462 | -20.7925
12.6114 | -34.6466 | -20.801
12.6175 | -34.6471 | -20.8097
12.6236 | -34.6475 | -20.8183
12.6297 | -34.648 | -20.8268
12.6358 | -34.6484 | -20.8352
12.6419 | -34.6488 | -20.8438
12.648 | -34.6493 | -20.8524
12.6541 | -34.6497 | -20.861
12.6602 | -34.6502 | -20.8692
12.6663 | -34.6506 | -20.8778
12.6724 | -34.6511 | -20.8863
12.6785 | -34.6516 | -20.8948
12.6847 | -34.652 | -20.9033
12.6908 | -34.6525 | -20.9118
12.6969 | -34.6529 | -20.9203
12.703 | -34.6534 | -20.9289
12.7091 | -34.6538 | -20.9375
12.7152 | -34.6543 | -20.946
12.7213 | -34.6547 | -20.9546
12.7274 | -34.6551 | -20.9632
12.7335 | -34.6556 | -20.9714
12.7396 | -34.656 | -20.9799
12.7457 | -34.6565 | -20.9885
12.7518 | -34.6569 | -20.997
12.7579 | -34.6574 | -21.0056
12.764 | -34.6578 | -21.0142
12.7701 | -34.6583 | -21.0227
12.7762 | -34.6587 | -21.0311
12.7823 | -34.6592 | -21.0397
12.7884 | -34.6597 | -21.0482
12.7945 | -34.6601 | -21.0568
12.8006 | -34.6606 | -21.0653
12.8067 | -34.661 | -21.0738
12.8128 | -34.6615 | -21.0822
12.8189 | -34.6619 | -21.0907
12.8251 | -34.6623 | -21.0992
12.8312 | -34.6628 | -21.1078
12.8373 | -34.6632 | -21.1163
12.8434 | -34.6637 | -21.1248
12.8495 | -34.6642 | -21.1332
12.8556 | -34.6646 | -21.1416
12.8617 | -34.665 | -21.1502
12.8678 | -34.6655 | -21.1588
12.8739 | -34.666 | -21.1672
12.88 | -34.6664 | -21.1757
12.8861 | -34.6668 | -21.1841
12.8922 | -34.6673 | -21.1927
12.8983 | -34.6677 | -21.2012
12.9044 | -34.6682 | -21.2095
12.9105 | -34.6686 | -21.218
12.9166 | -34.6691 | -21.2264
12.9227 | -34.6695 | -21.2349
12.9288 | -34.67 | -21.2434
12.9349 | -34.6704 | -21.252
12.941 | -34.6709 | -21.2604
12.9471 | -34.6714 | -21.2689
12.9532 | -34.6719 | -21.2774
12.9593 | -34.6723 | -21.2861
12.9654 | -34.6728 | -21.2945
12.9716 | -34.6732 | -21.303
12.9777 | -34.6737 | -21.3116
12.9838 | -34.6742 | -21.32
12.9899 | -34.6746 | -21.3286
12.996 | -34.675 | -21.3371
13.0021 | -34.6755 | -21.3456
13.0082 | -34.676 | -21.354
13.0143 | -34.6765 | -21.3623
13.0204 | -34.6769 | -21.3707
13.0265 | -34.6774 | -21.3792
13.0326 | -34.6778 | -21.3877
13.0387 | -34.6783 | -21.3964
13.0448 | -34.6787 | -21.4048
13.0509 | -34.6792 | -21.4133
13.057 | -34.6797 | -21.4216
13.0631 | -34.6801 | -21.4301
13.0692 | -34.6806 | -21.4385
13.0753 | -34.6811 | -21.4471
13.0814 | -34.6815 | -21.4557
13.0875 | -34.682 | -21.4639
13.0936 | -34.6824 | -21.4725
13.0997 | -34.6829 | -21.4809
13.1058 | -34.6833 | -21.4894
13.112 | -34.6838 | -21.4979
13.1181 | -34.6843 | -21.5063
13.1242 | -34.6847 | -21.5149
13.1303 | -34.6852 | -21.5234
13.1364 | -34.6856 | -21.5317
13.1425 | -34.686 | -21.5402
13.1486 | -34.6865 | -21.5488
13.1547 | -34.687 | -21.5573
13.1608 | -34.6874 | -21.5659
13.1669 | -34.6879 | -21.5742
13.173 | -34.6884 | -21.5827
13.1791 | -34.6888 | -21.5911
13.1852 | -34.6893 | -21.5997
13.1913 | -34.6897 | -21.6081
13.1974 | -34.6902 | -21.6166
13.2035 | -34.6906 | -21.6251
13.2096 | -34.6911 | -21.6335
13.2157 | -34.6916 | -21.642
13.2218 | -34.692 | -21.6504
13.2279 | -34.6925 | -21.6588
13.234 | -34.6929 | -21.6672
13.2401 | -34.6934 | -21.6757
13.2462 | -34.6938 | -21.6842
13.2524 | -34.6943 | -21.6927
13.2585 | -34.6948 | -21.7012
13.2646 | -34.6952 | -21.7098
13.2707 | -34.6957 | -21.7182
13.2768 | -34.6961 | -21.7268
13.2829 | -34.6966 | -21.7354
13.289 | -34.6971 | -21.7437
13.2951 | -34.6975 | -21.7521
13.3012 | -34.698 | -21.7605
13.3073 | -34.6984 | -21.7689
13.3134 | -34.6989 | -21.7774
13.3195 | -34.6993 | -21.7858
13.3256 | -34.6998 | -21.7942
13.3317 | -34.7002 | -21.8027
13.3378 | -34.7007 | -21.8113
13.3439 | -34.7011 | -21.8198
13.35 | -34.7016 | -21.8284
13.3561 | -34.702 | -21.8367
13.3622 | -34.7025 | -21.8451
13.3683 | -34.703 | -21.8535
13.3744 | -34.7034 | -21.862
13.3805 | -34.7039 | -21.8704
13.3866 | -34.7043 | -21.8789
13.3927 | -34.7048 | -21.8876
13.3989 | -34.7052 | -21.896
13.405 | -34.7057 | -21.9046
13.4111 | -34.7062 | -21.913
13.4172 | -34.7066 | -21.9216
13.4233 | -34.7071 | -21.9302
13.4294 | -34.7075 | -21.9387
13.4355 | -34.708 | -21.9471
13.4416 | -34.7084 | -21.9555
13.4477 | -34.7089 | -21.9639
13.4538 | -34.7094 | -21.9724
13.4599 | -34.7098 | -21.9809
13.466 | -34.7102 | -21.9894
13.4721 | -34.7107 | -21.9977
13.4782 | -34.7111 | -22.0062
13.4843 | -34.7116 | -22.0148
13.4904 | -34.7121 | -22.0231
13.4965 | -34.7125 | -22.0316
13.5026 | -34.713 | -22.0402
13.5087 | -34.7134 | -22.0486
13.5148 | -34.7139 | -22.0571
13.5209 | -34.7143 | -22.0657
13.527 | -34.7148 | -22.0742
13.5331 | -34.7153 | -22.0826
13.5393 | -34.7157 | -22.0911
13.5454 | -34.7162 | -22.0994
13.5515 | -34.7166 | -22.1081
13.5576 | -34.7171 | -22.1165
13.5637 | -34.7176 | -22.1251
13.5698 | -34.718 | -22.1336
13.5759 | -34.7185 | -22.1416
13.582 | -34.7189 | -22.15
13.5881 | -34.7194 | -22.1584
13.5942 | -34.7198 | -22.167
13.6003 | -34.7203 | -22.1754
13.6064 | -34.7208 | -22.184
13.6125 | -34.7212 | -22.1925
13.6186 | -34.7217 | -22.2009
13.6247 | -34.7222 | -22.2094
13.6308 | -34.7226 | -22.2178
13.6369 | -34.7231 | -22.2262
13.643 | -34.7235 | -22.2346
13.6491 | -34.724 | -22.2429
13.6552 | -34.7245 | -22.2514
13.6613 | -34.7249 | -22.2598
13.6674 | -34.7254 | -22.2682
13.6735 | -34.7258 | -22.2767
13.6796 | -34.7263 | -22.2851
13.6858 | -34.7267 | -22.2934
13.6919 | -34.7272 | -22.3018
13.698 | -34.7277 | -22.3103
13.7041 | -34.7281 | -22.319
13.7102 | -34.7286 | -22.3274
13.7163 | -34.7291 | -22.3358
13.7224 | -34.7295 | -22.3442
13.7285 | -34.73 | -22.3526
13.7346 | -34.7304 | -22.3612
13.7407 | -34.7309 | -22.3695
13.7468 | -34.7314 | -22.378
13.7529 | -34.7318 | -22.3863
13.759 | -34.7323 | -22.3948
13.7651 | -34.7327 | -22.4033
13.7712 | -34.7332 | -22.4118
13.7773 | -34.7336 | -22.4202
13.7834 | -34.7341 | -22.4286
13.7895 | -34.7345 | -22.4369
13.7956 | -34.735 | -22.4454
13.8017 | -34.7354 | -22.4537
13.8078 | -34.7359 | -22.4622
13.8139 | -34.7363 | -22.4706
13.82 | -34.7368 | -22.479
13.8262 | -34.7373 | -22.4876
13.8323 | -34.7378 | -22.4962
13.8384 | -34.7382 | -22.5046
13.8445 | -34.7386 | -22.513
13.8506 | -34.7391 | -22.5214
13.8567 | -34.7396 | -22.5299
13.8628 | -34.74 | -22.5384
13.8689 | -34.7405 | -22.5468
13.875 | -34.741 | -22.5553
13.8811 | -34.7414 | -22.5638
13.8872 | -34.7418 | -22.5722
13.8933 | -34.7423 | -22.5805
13.8994 | -34.7427 | -22.589
13.9055 | -34.7432 | -22.5975
13.9116 | -34.7436 | -22.6059
13.9177 | -34.7441 | -22.6143
13.9238 | -34.7445 | -22.6228
13.9299 | -34.745 | -22.6312
13.936 | -34.7454 | -22.6397
13.9421 | -34.7459 | -22.6481
13.9482 | -34.7463 | -22.6565
13.9543 | -34.7468 | -22.665
13.9604 | -34.7472 | -22.6735
13.9665 | -34.7477 | -22.6818
13.9727 | -34.7482 | -22.6904
13.9788 | -34.7487 | -22.6987
13.9849 | -34.7491 | -22.7072
13.991 | -34.7496 | -22.7157
13.9971 | -34.75 | -22.7241
14.0032 | -34.7505 | -22.7325
14.0093 | -34.7509 | -22.7409
14.0154 | -34.7514 | -22.7493
14.0215 | -34.7518 | -22.7576
14.0276 | -34.7523 | -22.766
14.0337 | -34.7528 | -22.7744
14.0398 | -34.7532 | -22.7829
14.0459 | -34.7537 | -22.7914
14.052 | -34.7542 | -22.7999
14.0581 | -34.7546 | -22.8084
14.0642 | -34.755 | -22.8169
14.0703 | -34.7555 | -22.8252
14.0764 | -34.756 | -22.8336
14.0825 | -34.7564 | -22.8421
14.0886 | -34.7569 | -22.8505
14.0947 | -34.7573 | -22.8589
14.1008 | -34.7578 | -22.8673
14.1069 | -34.7582 | -22.8756
14.1131 | -34.7587 | -22.8841
14.1192 | -34.7591 | -22.8926
14.1253 | -34.7596 | -22.9009
14.1314 | -34.7601 | -22.9092
14.1375 | -34.7605 | -22.9175
14.1436 | -34.761 | -22.9258
14.1497 | -34.7614 | -22.9343
14.1558 | -34.7619 | -22.9428
14.1619 | -34.7623 | -22.9512
14.168 | -34.7628 | -22.9596
14.1741 | -34.7633 | -22.9679
14.1802 | -34.7637 | -22.9763
14.1863 | -34.7642 | -22.9848
14.1924 | -34.7647 | -22.9932
14.1985 | -34.7651 | -23.0016
14.2046 | -34.7656 | -23.0101
14.2107 | -34.7661 | -23.0183
14.2168 | -34.7665 | -23.0267
14.2229 | -34.767 | -23.0351
14.229 | -34.7674 | -23.0435
14.2351 | -34.7679 | -23.0518
14.2412 | -34.7683 | -23.0603
14.2473 | -34.7688 | -23.0686
14.2534 | -34.7692 | -23.0768
14.2596 | -34.7697 | -23.0851
14.2657 | -34.7701 | -23.0935
14.2718 | -34.7706 | -23.102
14.2779 | -34.7711 | -23.1104
14.284 | -34.7715 | -23.1188
14.2901 | -34.772 | -23.1271
14.2962 | -34.7725 | -23.1355
14.3023 | -34.7729 | -23.144
14.3084 | -34.7734 | -23.1523
14.3145 | -34.7738 | -23.1605
14.3206 | -34.7743 | -23.1691
14.3267 | -34.7747 | -23.1775
14.3328 | -34.7752 | -23.1858
14.3389 | -34.7757 | -23.1942
14.345 | -34.7761 | -23.2025
14.3511 | -34.7766 | -23.2109
14.3572 | -34.777 | -23.2194
14.3633 | -34.7775 | -23.2279
14.3694 | -34.778 | -23.2362
14.3755 | -34.7784 | -23.2446
14.3816 | -34.7789 | -23.2531
14.3877 | -34.7794 | -23.2613
14.3938 | -34.7798 | -23.2698
14.4 | -34.7803 | -23.2782
14.4061 | -34.7807 | -23.2865
14.4122 | -34.7812 | -23.2949
14.4183 | -34.7817 | -23.3033
14.4244 | -34.7821 | -23.3116
14.4305 | -34.7826 | -23.3199
14.4366 | -34.7831 | -23.3283
14.4427 | -34.7835 | -23.3368
14.4488 | -34.784 | -23.3452
14.4549 | -34.7845 | -23.3534
14.461 | -34.7849 | -23.3617
14.4671 | -34.7854 | -23.37
14.4732 | -34.7858 | -23.3784
14.4793 | -34.7863 | -23.3867
14.4854 | -34.7868 | -23.395
14.4915 | -34.7873 | -23.4031
14.4976 | -34.7877 | -23.4113
14.5037 | -34.7882 | -23.4196
14.5098 | -34.7887 | -23.4279
14.5159 | -34.7891 | -23.4363
14.522 | -34.7896 | -23.4445
14.5281 | -34.79 | -23.4526
14.5342 | -34.7905 | -23.461
14.5403 | -34.791 | -23.4692
14.5465 | -34.7914 | -23.4776
14.5526 | -34.7919 | -23.4861
14.5587 | -34.7923 | -23.4946
14.5648 | -34.7928 | -23.5027
14.5709 | -34.7933 | -23.5109
14.577 | -34.7937 | -23.5192
14.5831 | -34.7941 | -23.5275
14.5892 | -34.7946 | -23.5358
14.5953 | -34.7951 | -23.5439
14.6014 | -34.7955 | -23.5522
14.6075 | -34.796 | -23.5604
14.6136 | -34.7964 | -23.5688
14.6197 | -34.7969 | -23.5772
14.6258 | -34.7973 | -23.5856
14.6319 | -34.7978 | -23.5938
14.638 | -34.7982 | -23.602
14.6441 | -34.7987 | -23.6105
14.6502 | -34.7992 | -23.619
14.6563 | -34.7996 | -23.6274
14.6624 | -34.8001 | -23.6357
14.6685 | -34.8006 | -23.6437
14.6746 | -34.801 | -23.6522
14.6807 | -34.8015 | -23.6603
14.6869 | -34.802 | -23.6686
14.693 | -34.8024 | -23.6769
14.6991 | -34.8029 | -23.6851
14.7052 | -34.8034 | -23.6932
14.7113 | -34.8038 | -23.7016
14.7174 | -34.8043 | -23.7099
14.7235 | -34.8048 | -23.7181
14.7296 | -34.8052 | -23.7263
14.7357 | -34.8057 | -23.7345
14.7418 | -34.8062 | -23.7428
14.7479 | -34.8067 | -23.751
14.754 | -34.8071 | -23.7593
14.7601 | -34.8076 | -23.7675
14.7662 | -34.8081 | -23.7758
14.7723 | -34.8085 | -23.7837
14.7784 | -34.809 | -23.792
14.7845 | -34.8094 | -23.8003
14.7906 | -34.8099 | -23.8086
14.7967 | -34.8104 | -23.8168
14.8028 | -34.8108 | -23.825
14.8089 | -34.8113 | -23.8332
14.815 | -34.8118 | -23.8415
14.8211 | -34.8122 | -23.8498
14.8272 | -34.8127 | -23.8581
14.8334 | -34.8132 | -23.8663
14.8395 | -34.8136 | -23.8745
14.8456 | -34.8141 | -23.8826
14.8517 | -34.8146 | -23.8909
14.8578 | -34.815 | -23.8992
14.8639 | -34.8155 | -23.9075
14.87 | -34.816 | -23.9158
14.8761 | -34.8165 | -23.9238
14.8822 | -34.8169 | -23.932
14.8883 | -34.8174 | -23.9403
14.8944 | -34.8178 | -23.9484
14.9005 | -34.8183 | -23.9566
14.9066 | -34.8188 | -23.9648
14.9127 | -34.8192 | -23.9729
14.9188 | -34.8197 | -23.9811
14.9249 | -34.8201 | -23.9894
14.931 | -34.8206 | -23.9977
14.9371 | -34.8211 | -24.0058
14.9432 | -34.8215 | -24.014
14.9493 | -34.822 | -24.0221
14.9554 | -34.8225 | -24.0303
14.9615 | -34.823 | -24.0384
14.9676 | -34.8234 | -24.0468
14.9738 | -34.8239 | -24.0548
14.9799 | -34.8243 | -24.0632
14.986 | -34.8248 | -24.0711
14.9921 | -34.8253 | -24.0793
14.9982 | -34.8257 | -24.0874
15.0043 | -34.8262 | -24.0955
15.0104 | -34.8266 | -24.1038
15.0165 | -34.8271 | -24.1121
15.0226 | -34.8276 | -24.1201
15.0287 | -34.8281 | -24.1284
15.0348 | -34.8285 | -24.1365
15.0409 | -34.829 | -24.1448
15.047 | -34.8295 | -24.1531
15.0531 | -34.83 | -24.161
15.0592 | -34.8305 | -24.1692
15.0653 | -34.831 | -24.1773
15.0714 | -34.8314 | -24.1854
15.0775 | -34.8319 | -24.1935
15.0836 | -34.8323 | -24.2016
15.0897 | -34.8328 | -24.2094
15.0958 | -34.8333 | -24.2175
15.1019 | -34.8338 | -24.2256
15.108 | -34.8342 | -24.2337
15.1141 | -34.8347 | -24.2419
15.1203 | -34.8352 | -24.2501
15.1264 | -34.8357 | -24.2581
15.1325 | -34.8361 | -24.2662
15.1386 | -34.8366 | -24.2743
15.1447 | -34.8371 | -24.2824
15.1508 | -34.8375 | -24.2905
15.1569 | -34.838 | -24.2985
15.163 | -34.8385 | -24.3063
15.1691 | -34.839 | -24.3143
15.1752 | -34.8394 | -24.3224
15.1813 | -34.8399 | -24.3306
15.1874 | -34.8404 | -24.3388
15.1935 | -34.8409 | -24.3466
15.1996 | -34.8413 | -24.3546
15.2057 | -34.8418 | -24.3628
15.2118 | -34.8423 | -24.371
15.2179 | -34.8428 | -24.3793
15.224 | -34.8432 | -24.3873
15.2301 | -34.8437 | -24.3952
15.2362 | -34.8442 | -24.4033
15.2423 | -34.8447 | -24.4116
15.2484 | -34.8452 | -24.4196
15.2545 | -34.8456 | -24.4277
15.2607 | -34.8461 | -24.4358
15.2668 | -34.8466 | -24.4435
15.2729 | -34.8471 | -24.4515
15.279 | -34.8475 | -24.4597
15.2851 | -34.848 | -24.4677
15.2912 | -34.8485 | -24.4758
15.2973 | -34.849 | -24.4839
15.3034 | -34.8494 | -24.4918
15.3095 | -34.8499 | -24.4998
15.3156 | -34.8504 | -24.5078
15.3217 | -34.8509 | -24.5158
15.3278 | -34.8513 | -24.524
15.3339 | -34.8518 | -24.5318
15.34 | -34.8523 | -24.5397
15.3461 | -34.8528 | -24.5477
15.3522 | -34.8533 | -24.5556
15.3583 | -34.8537 | -24.5638
15.3644 | -34.8542 | -24.5717
15.3705 | -34.8547 | -24.5794
15.3766 | -34.8552 | -24.5875
15.3827 | -34.8557 | -24.5954
15.3888 | -34.8561 | -24.6035
15.3949 | -34.8566 | -24.6116
15.401 | -34.8571 | -24.6197
15.4072 | -34.8575 | -24.6275
15.4133 | -34.858 | -24.6355
15.4194 | -34.8585 | -24.6435
15.4255 | -34.8589 | -24.6516
15.4316 | -34.8594 | -24.6595
15.4377 | -34.8599 | -24.6676
15.4438 | -34.8604 | -24.6754
15.4499 | -34.8609 | -24.6834
15.456 | -34.8613 | -24.6913
15.4621 | -34.8618 | -24.6993
15.4682 | -34.8623 | -24.7074
15.4743 | -34.8628 | -24.7151
15.4804 | -34.8632 | -24.7231
15.4865 | -34.8637 | -24.7311
15.4926 | -34.8642 | -24.7391
15.4987 | -34.8646 | -24.747
15.5048 | -34.8651 | -24.7551
15.5109 | -34.8656 | -24.763
15.517 | -34.8661 | -24.771
15.5231 | -34.8666 | -24.7791
15.5292 | -34.8671 | -24.7871
15.5353 | -34.8675 | -24.7951
15.5414 | -34.868 | -24.803
15.5476 | -34.8685 | -24.8105
15.5537 | -34.869 | -24.8185
15.5598 | -34.8695 | -24.8267
15.5659 | -34.8699 | -24.8346
15.572 | -34.8704 | -24.8424
15.5781 | -34.8709 | -24.8504
15.5842 | -34.8714 | -24.8581
15.5903 | -34.8719 | -24.8659
15.5964 | -34.8724 | -24.8738
15.6025 | -34.8728 | -24.8817
15.6086 | -34.8733 | -24.8897
15.6147 | -34.8738 | -24.8978
15.6208 | -34.8743 | -24.9055
15.6269 | -34.8748 | -24.9134
15.633 | -34.8753 | -24.9213
15.6391 | -34.8757 | -24.9292
15.6452 | -34.8762 | -24.9371
15.6513 | -34.8767 | -24.9449
15.6574 | -34.8772 | -24.9529
15.6635 | -34.8777 | -24.9608
15.6696 | -34.8782 | -24.9689
15.6757 | -34.8787 | -24.9768
15.6818 | -34.8791 | -24.9845
15.688 | -34.8796 | -24.9923
15.6941 | -34.8801 | -25.0002
15.7002 | -34.8805 | -25.008
15.7063 | -34.881 | -25.0159
15.7124 | -34.8815 | -25.0239
15.7185 | -34.8819 | -25.0319
15.7246 | -34.8825 | -25.0395
15.7307 | -34.8829 | -25.0473
15.7368 | -34.8834 | -25.0552
15.7429 | -34.8839 | -25.0629
15.749 | -34.8844 | -25.0707
15.7551 | -34.8849 | -25.0786
15.7612 | -34.8854 | -25.0862
15.7673 | -34.8858 | -25.0942
15.7734 | -34.8863 | -25.1019
15.7795 | -34.8868 | -25.1098
15.7856 | -34.8873 | -25.1176
15.7917 | -34.8878 | -25.1252
15.7978 | -34.8882 | -25.1332
15.8039 | -34.8887 | -25.141
15.81 | -34.8892 | -25.1489
15.8161 | -34.8897 | -25.1568
15.8222 | -34.8902 | -25.1645
15.8283 | -34.8907 | -25.172
15.8345 | -34.8912 | -25.1799
15.8406 | -34.8917 | -25.1879
15.8467 | -34.8922 | -25.1956
15.8528 | -34.8927 | -25.2035
15.8589 | -34.8932 | -25.2115
15.865 | -34.8937 | -25.2191
15.8711 | -34.8941 | -25.227
15.8772 | -34.8946 | -25.2347
15.8833 | -34.8951 | -25.2425
15.8894 | -34.8956 | -25.2505
15.8955 | -34.8961 | -25.2582
15.9016 | -34.8965 | -25.2657
15.9077 | -34.897 | -25.2735
15.9138 | -34.8975 | -25.2814
15.9199 | -34.898 | -25.2891
15.926 | -34.8985 | -25.2968
15.9321 | -34.899 | -25.3045
15.9382 | -34.8994 | -25.3123
15.9443 | -34.8999 | -25.32
15.9504 | -34.9004 | -25.3278
15.9565 | -34.9009 | -25.3356
15.9626 | -34.9013 | -25.3435
15.9687 | -34.9018 | -25.351
15.9749 | -34.9023 | -25.3586
15.981 | -34.9028 | -25.3664
15.9871 | -34.9032 | -25.3744
15.9932 | -34.9037 | -25.3821
15.9993 | -34.9042 | -25.3897
16.0054 | -34.9047 | -25.3971
16.0115 | -34.9052 | -25.405
16.0176 | -34.9056 | -25.4127
16.0237 | -34.9061 | -25.4205
16.0298 | -34.9066 | -25.4282
16.0359 | -34.907 | -25.436
16.042 | -34.9076 | -25.4437
16.0481 | -34.908 | -25.4516
16.0542 | -34.9085 | -25.4594
16.0603 | -34.909 | -25.4669
16.0664 | -34.9094 | -25.4746
16.0725 | -34.9099 | -25.482
16.0786 | -34.9104 | -25.4897
16.0847 | -34.9109 | -25.4975
16.0908 | -34.9114 | -25.5052
16.0969 | -34.9119 | -25.513
16.103 | -34.9124 | -25.5208
16.1091 | -34.9129 | -25.5283
16.1152 | -34.9134 | -25.536
16.1214 | -34.9139 | -25.5438
16.1275 | -34.9143 | -25.5516
16.1336 | -34.9148 | -25.5593
16.1397 | -34.9153 | -25.567
16.1458 | -34.9158 | -25.5745
16.1519 | -34.9163 | -25.5822
16.158 | -34.9168 | -25.5901
16.1641 | -34.9172 | -25.5979
16.1702 | -34.9177 | -25.6056
16.1763 | -34.9182 | -25.6133
16.1824 | -34.9187 | -25.6208
16.1885 | -34.9192 | -25.6284
16.1946 | -34.9197 | -25.636
16.2007 | -34.9201 | -25.6439
16.2068 | -34.9206 | -25.6514
16.2129 | -34.9211 | -25.6592
16.219 | -34.9216 | -25.6665
16.2251 | -34.9221 | -25.6742
16.2312 | -34.9226 | -25.6819
16.2373 | -34.9231 | -25.6896
16.2434 | -34.9236 | -25.6973
16.2495 | -34.9241 | -25.7047
16.2556 | -34.9246 | -25.7126
16.2618 | -34.9251 | -25.7203
16.2679 | -34.9255 | -25.728
16.274 | -34.926 | -25.7356
16.2801 | -34.9265 | -25.7432
16.2862 | -34.927 | -25.7508
16.2923 | -34.9275 | -25.7585
16.2984 | -34.9279 | -25.7661
16.3045 | -34.9284 | -25.7737
16.3106 | -34.9289 | -25.7814
16.3167 | -34.9294 | -25.7891
16.3228 | -34.9299 | -25.7965
16.3289 | -34.9304 | -25.8041
16.335 | -34.9308 | -25.8116
16.3411 | -34.9313 | -25.8193
16.3472 | -34.9318 | -25.8271
16.3533 | -34.9323 | -25.8347
16.3594 | -34.9328 | -25.8422
16.3655 | -34.9333 | -25.8499
16.3716 | -34.9338 | -25.8572
16.3777 | -34.9343 | -25.8648
16.3838 | -34.9347 | -25.8724
16.3899 | -34.9352 | -25.8799
16.396 | -34.9357 | -25.8876
16.4021 | -34.9362 | -25.8952
16.4083 | -34.9367 | -25.9027
16.4144 | -34.9372 | -25.9104
16.4205 | -34.9377 | -25.9181
16.4266 | -34.9382 | -25.9255
16.4327 | -34.9386 | -25.9332
16.4388 | -34.9391 | -25.9409
16.4449 | -34.9396 | -25.9485
16.451 | -34.9401 | -25.9562
16.4571 | -34.9406 | -25.964
16.4632 | -34.9411 | -25.9711
16.4693 | -34.9415 | -25.9787
16.4754 | -34.942 | -25.9864
16.4815 | -34.9425 | -25.994
16.4876 | -34.943 | -26.0014
16.4937 | -34.9434 | -26.009
16.4998 | -34.9439 | -26.0163
16.5059 | -34.9444 | -26.0237
16.512 | -34.9449 | -26.0315
16.5181 | -34.9454 | -26.0393
16.5242 | -34.9459 | -26.0471
16.5303 | -34.9464 | -26.0546
16.5364 | -34.9468 | -26.0621
16.5425 | -34.9473 | -26.0697
16.5487 | -34.9478 | -26.0773
16.5548 | -34.9483 | -26.0849
16.5609 | -34.9488 | -26.0925
16.567 | -34.9493 | -26.0999
16.5731 | -34.9498 | -26.1075
16.5792 | -34.9502 | -26.1151
16.5853 | -34.9507 | -26.1227
16.5914 | -34.9512 | -26.1304
16.5975 | -34.9517 | -26.138
16.6036 | -34.9521 | -26.1455
16.6097 | -34.9526 | -26.1531
16.6158 | -34.9531 | -26.1606
16.6219 | -34.9536 | -26.1681
16.628 | -34.9541 | -26.1757
16.6341 | -34.9545 | -26.1833
16.6402 | -34.955 | -26.1906
16.6463 | -34.9555 | -26.1984
16.6524 | -34.956 | -26.2059
16.6585 | -34.9565 | -26.2136
16.6646 | -34.957 | -26.2211
16.6707 | -34.9575 | -26.2285
16.6768 | -34.958 | -26.2361
16.6829 | -34.9585 | -26.2435
16.689 | -34.959 | -26.2509
16.6952 | -34.9594 | -26.2586
16.7013 | -34.9599 | -26.266
16.7074 | -34.9604 | -26.2733
16.7135 | -34.9609 | -26.2809
16.7196 | -34.9614 | -26.2884
16.7257 | -34.9618 | -26.2959
16.7318 | -34.9623 | -26.3034
16.7379 | -34.9628 | -26.311
16.744 | -34.9633 | -26.3183
16.7501 | -34.9638 | -26.3259
16.7562 | -34.9642 | -26.3335
16.7623 | -34.9647 | -26.341
16.7684 | -34.9652 | -26.3486
16.7745 | -34.9657 | -26.3563
16.7806 | -34.9662 | -26.3636
16.7867 | -34.9667 | -26.3711
16.7928 | -34.9672 | -26.3786
16.7989 | -34.9676 | -26.3861
16.805 | -34.9681 | -26.3937
16.8111 | -34.9686 | -26.401
16.8172 | -34.9691 | -26.4086
16.8233 | -34.9696 | -26.4162
16.8294 | -34.9701 | -26.4237
16.8356 | -34.9705 | -26.4313
16.8417 | -34.971 | -26.4387
16.8478 | -34.9715 | -26.446
16.8539 | -34.972 | -26.4534
16.86 | -34.9725 | -26.461
16.8661 | -34.9729 | -26.4685
16.8722 | -34.9735 | -26.476
16.8783 | -34.9739 | -26.4835
16.8844 | -34.9744 | -26.4909
16.8905 | -34.9749 | -26.4984
16.8966 | -34.9754 | -26.5059
16.9027 | -34.9759 | -26.5134
16.9088 | -34.9764 | -26.521
16.9149 | -34.9769 | -26.5286
16.921 | -34.9774 | -26.536
16.9271 | -34.9779 | -26.5437
16.9332 | -34.9783 | -26.5512
16.9393 | -34.9788 | -26.5588
16.9454 | -34.9793 | -26.5662
16.9515 | -34.9798 | -26.5736
16.9576 | -34.9803 | -26.5808
16.9637 | -34.9808 | -26.5883
16.9698 | -34.9813 | -26.5957
16.9759 | -34.9817 | -26.6033
16.9821 | -34.9822 | -26.6109
16.9882 | -34.9828 | -26.6183
16.9943 | -34.9832 | -26.6257
17.0004 | -34.9837 | -26.6332
17.0065 | -34.9842 | -26.6407
17.0126 | -34.9847 | -26.6482
17.0187 | -34.9852 | -26.6557
17.0248 | -34.9857 | -26.6627
17.0309 | -34.9862 | -26.6702
17.037 | -34.9867 | -26.6776
17.0431 | -34.9872 | -26.6851
17.0492 | -34.9877 | -26.6925
17.0553 | -34.9882 | -26.6999
17.0614 | -34.9887 | -26.7074
17.0675 | -34.9892 | -26.7149
17.0736 | -34.9897 | -26.7225
17.0797 | -34.9901 | -26.7299
17.0858 | -34.9906 | -26.7374
17.0919 | -34.9911 | -26.7447
17.098 | -34.9916 | -26.752
17.1041 | -34.9921 | -26.7594
17.1102 | -34.9926 | -26.7667
17.1163 | -34.9931 | -26.774
17.1225 | -34.9935 | -26.7813
17.1286 | -34.9941 | -26.7886
17.1347 | -34.9945 | -26.796
17.1408 | -34.995 | -26.8034
17.1469 | -34.9955 | -26.8108
17.153 | -34.996 | -26.8182
17.1591 | -34.9965 | -26.8257
17.1652 | -34.997 | -26.833
17.1713 | -34.9975 | -26.8405
17.1774 | -34.998 | -26.8478
17.1835 | -34.9984 | -26.8552
17.1896 | -34.9989 | -26.8625
17.1957 | -34.9994 | -26.8699
17.2018 | -34.9999 | -26.8773
17.2079 | -35.0004 | -26.8847
17.214 | -35.0008 | -26.8922
17.2201 | -35.0013 | -26.8995
17.2262 | -35.0018 | -26.907
17.2323 | -35.0023 | -26.9144
17.2384 | -35.0028 | -26.9218
17.2445 | -35.0032 | -26.9292
17.2506 | -35.0037 | -26.9366
17.2567 | -35.0042 | -26.944
17.2628 | -35.0047 | -26.9515
17.269 | -35.0052 | -26.9586
17.2751 | -35.0057 | -26.9659
17.2812 | -35.0061 | -26.9735
17.2873 | -35.0067 | -26.981
17.2934 | -35.0071 | -26.9884
17.2995 | -35.0076 | -26.9958
17.3056 | -35.0081 | -27.0029
17.3117 | -35.0086 | -27.0104
17.3178 | -35.0091 | -27.0177
17.3239 | -35.0096 | -27.025
17.33 | -35.01 | -27.0325
17.3361 | -35.0105 | -27.0399
17.3422 | -35.011 | -27.0469
17.3483 | -35.0115 | -27.0543
17.3544 | -35.012 | -27.0618
17.3605 | -35.0125 | -27.0691
17.3666 | -35.0129 | -27.0765
17.3727 | -35.0134 | -27.0839
17.3788 | -35.0139 | -27.0911
17.3849 | -35.0144 | -27.0983
17.391 | -35.0149 | -27.1057
17.3971 | -35.0154 | -27.1131
17.4032 | -35.0159 | -27.1204
17.4094 | -35.0163 | -27.1276
17.4155 | -35.0168 | -27.1349
17.4216 | -35.0173 | -27.1421
17.4277 | -35.0178 | -27.1496
17.4338 | -35.0183 | -27.1569
17.4399 | -35.0188 | -27.1644
17.446 | -35.0193 | -27.1716
17.4521 | -35.0197 | -27.1789
17.4582 | -35.0202 | -27.1862
17.4643 | -35.0207 | -27.1935
17.4704 | -35.0212 | -27.2009
17.4765 | -35.0217 | -27.2084
17.4826 | -35.0221 | -27.2156
17.4887 | -35.0226 | -27.223
17.4948 | -35.0231 | -27.2305
17.5009 | -35.0236 | -27.2379
17.507 | -35.024 | -27.2453
17.5131 | -35.0246 | -27.2526
17.5192 | -35.0251 | -27.2598
17.5253 | -35.0256 | -27.2671
17.5314 | -35.026 | -27.2744
17.5375 | -35.0265 | -27.2819
17.5436 | -35.027 | -27.2892
17.5497 | -35.0275 | -27.2965
17.5559 | -35.028 | -27.3035
17.562 | -35.0285 | -27.3109
17.5681 | -35.029 | -27.3182
17.5742 | -35.0295 | -27.3257
17.5803 | -35.0299 | -27.333
17.5864 | -35.0304 | -27.3402
17.5925 | -35.0309 | -27.3476
17.5986 | -35.0314 | -27.3548
17.6047 | -35.0319 | -27.3621
17.6108 | -35.0324 | -27.3693
17.6169 | -35.0329 | -27.3767
17.623 | -35.0334 | -27.3837
17.6291 | -35.0339 | -27.391
17.6352 | -35.0344 | -27.3983
17.6413 | -35.0348 | -27.4057
17.6474 | -35.0353 | -27.4131
17.6535 | -35.0358 | -27.4205
17.6596 | -35.0363 | -27.4277
17.6657 | -35.0368 | -27.4351
17.6718 | -35.0373 | -27.4425
17.6779 | -35.0377 | -27.4499
17.684 | -35.0382 | -27.4573
17.6901 | -35.0387 | -27.4648
17.6963 | -35.0392 | -27.4719
17.7024 | -35.0397 | -27.4791
17.7085 | -35.0402 | -27.4863
17.7146 | -35.0407 | -27.4936
17.7207 | -35.0412 | -27.501
17.7268 | -35.0417 | -27.5081
17.7329 | -35.0421 | -27.5156
17.739 | -35.0426 | -27.523
17.7451 | -35.0431 | -27.5304
17.7512 | -35.0436 | -27.5377
17.7573 | -35.0441 | -27.545
17.7634 | -35.0446 | -27.552
17.7695 | -35.0451 | -27.5592
17.7756 | -35.0456 | -27.5666
17.7817 | -35.0461 | -27.5739
17.7878 | -35.0466 | -27.5814
17.7939 | -35.0471 | -27.5887
17.8 | -35.0475 | -27.5961
17.8061 | -35.048 | -27.6033
17.8122 | -35.0486 | -27.6103
17.8183 | -35.0491 | -27.6175
17.8244 | -35.0495 | -27.6247
17.8305 | -35.05 | -27.632
17.8366 | -35.0505 | -27.639
17.8428 | -35.051 | -27.6462
17.8489 | -35.0515 | -27.6534
17.855 | -35.052 | -27.6605
17.8611 | -35.0525 | -27.6677
17.8672 | -35.0529 | -27.6749
17.8733 | -35.0534 | -27.6823
17.8794 | -35.0539 | -27.6894
17.8855 | -35.0544 | -27.6967
17.8916 | -35.0548 | -27.704
17.8977 | -35.0553 | -27.7114
17.9038 | -35.0558 | -27.7181
17.9099 | -35.0563 | -27.7252
17.916 | -35.0568 | -27.7326
17.9221 | -35.0573 | -27.7399
17.9282 | -35.0578 | -27.747
17.9343 | -35.0583 | -27.7543
17.9404 | -35.0588 | -27.7615
17.9465 | -35.0593 | -27.7687
17.9526 | -35.0597 | -27.7759
17.9587 | -35.0602 | -27.7835
17.9648 | -35.0607 | -27.7906
17.9709 | -35.0612 | -27.7981
17.977 | -35.0617 | -27.8052
17.9832 | -35.0622 | -27.8125
17.9893 | -35.0626 | -27.8197
17.9954 | -35.0631 | -27.827
18.0015 | -35.0636 | -27.8343
18.0076 | -35.0641 | -27.8414
18.0137 | -35.0645 | -27.8486
18.0198 | -35.065 | -27.8559
18.0259 | -35.0654 | -27.8632
18.032 | -35.0659 | -27.8705
18.0381 | -35.0664 | -27.8778
18.0442 | -35.0669 | -27.8849
18.0503 | -35.0674 | -27.8922
18.0564 | -35.0679 | -27.8994
18.0625 | -35.0683 | -27.9068
18.0686 | -35.0688 | -27.9141
18.0747 | -35.0693 | -27.9214
18.0808 | -35.0698 | -27.9284
18.0869 | -35.0703 | -27.9357
18.093 | -35.0708 | -27.9429
18.0991 | -35.0712 | -27.9501
18.1052 | -35.0717 | -27.9573
18.1113 | -35.0722 | -27.9646
18.1174 | -35.0727 | -27.9718
18.1236 | -35.0732 | -27.979
18.1297 | -35.0736 | -27.9861
18.1358 | -35.0741 | -27.9933
18.1419 | -35.0746 | -28.0003
18.148 | -35.0751 | -28.0077
18.1541 | -35.0756 | -28.0146
18.1602 | -35.076 | -28.0219
18.1663 | -35.0765 | -28.0291
18.1724 | -35.077 | -28.0365
18.1785 | -35.0775 | -28.0435
18.1846 | -35.0779 | -28.0506
18.1907 | -35.0784 | -28.0578
18.1968 | -35.0789 | -28.0651
18.2029 | -35.0794 | -28.0724
18.209 | -35.0799 | -28.0798
18.2151 | -35.0804 | -28.087
18.2212 | -35.0809 | -28.0942
18.2273 | -35.0813 | -28.1014
18.2334 | -35.0818 | -28.1086
18.2395 | -35.0823 | -28.1158
18.2456 | -35.0828 | -28.1231
18.2517 | -35.0833 | -28.1304
18.2578 | -35.0838 | -28.1375
18.2639 | -35.0843 | -28.1447
18.2701 | -35.0848 | -28.1519
18.2762 | -35.0853 | -28.1591
18.2823 | -35.0858 | -28.1663
18.2884 | -35.0862 | -28.1735
18.2945 | -35.0867 | -28.1802
18.3006 | -35.0872 | -28.1875
18.3067 | -35.0877 | -28.1947
18.3128 | -35.0882 | -28.2019
18.3189 | -35.0887 | -28.2091
18.325 | -35.0891 | -28.2163
18.3311 | -35.0896 | -28.2235
18.3372 | -35.0901 | -28.2307
18.3433 | -35.0906 | -28.2379
18.3494 | -35.091 | -28.245
18.3555 | -35.0915 | -28.2522
18.3616 | -35.092 | -28.2592
18.3677 | -35.0925 | -28.2665
18.3738 | -35.093 | -28.2736
18.3799 | -35.0935 | -28.2808
18.386 | -35.0939 | -28.2881
18.3921 | -35.0944 | -28.2952
18.3982 | -35.0949 | -28.3022
18.4043 | -35.0953 | -28.3094
18.4105 | -35.0958 | -28.3165
18.4166 | -35.0963 | -28.3238
18.4227 | -35.0968 | -28.3312
18.4288 | -35.0972 | -28.3384
18.4349 | -35.0977 | -28.3454
18.441 | -35.0981 | -28.3526
18.4471 | -35.0986 | -28.3597
18.4532 | -35.0991 | -28.367
18.4593 | -35.0996 | -28.3739
18.4654 | -35.1001 | -28.3809
18.4715 | -35.1006 | -28.388
18.4776 | -35.101 | -28.3954
18.4837 | -35.1015 | -28.4027
18.4898 | -35.102 | -28.4098
18.4959 | -35.1025 | -28.417
18.502 | -35.1029 | -28.4241
18.5081 | -35.1034 | -28.4314
18.5142 | -35.1039 | -28.4386
18.5203 | -35.1043 | -28.4457
18.5264 | -35.1048 | -28.4527
18.5325 | -35.1053 | -28.4601
18.5386 | -35.1058 | -28.4671
18.5447 | -35.1062 | -28.4742
18.5508 | -35.1067 | -28.4812
18.557 | -35.1072 | -28.4885
18.5631 | -35.1076 | -28.4957
18.5692 | -35.1081 | -28.5028
18.5753 | -35.1086 | -28.5097
18.5814 | -35.1091 | -28.5167
18.5875 | -35.1095 | -28.5239
18.5936 | -35.11 | -28.5311
18.5997 | -35.1105 | -28.5383
18.6058 | -35.1109 | -28.5453
18.6119 | -35.1114 | -28.5524
18.618 | -35.1119 | -28.5595
18.6241 | -35.1123 | -28.5667
18.6302 | -35.1128 | -28.5739
18.6363 | -35.1133 | -28.5809
18.6424 | -35.1138 | -28.5881
18.6485 | -35.1142 | -28.5954
18.6546 | -35.1147 | -28.6024
18.6607 | -35.1152 | -28.6096
18.6668 | -35.1156 | -28.6168
18.6729 | -35.1161 | -28.6239
18.679 | -35.1166 | -28.6307
18.6851 | -35.1171 | -28.6376
18.6912 | -35.1175 | -28.6447
18.6974 | -35.118 | -28.6518
18.7035 | -35.1184 | -28.659
18.7096 | -35.1189 | -28.6663
18.7157 | -35.1194 | -28.6734
18.7218 | -35.1199 | -28.6805
18.7279 | -35.1204 | -28.6876
18.734 | -35.1209 | -28.6948
18.7401 | -35.1214 | -28.7018
18.7462 | -35.1218 | -28.7086
18.7523 | -35.1223 | -28.7158
18.7584 | -35.1228 | -28.723
18.7645 | -35.1233 | -28.73
18.7706 | -35.1238 | -28.7371
18.7767 | -35.1243 | -28.7442
18.7828 | -35.1248 | -28.7514
18.7889 | -35.1252 | -28.7586
18.795 | -35.1257 | -28.766
18.8011 | -35.1262 | -28.773
18.8072 | -35.1267 | -28.7802
18.8133 | -35.1272 | -28.7873
18.8194 | -35.1277 | -28.7945
18.8255 | -35.1282 | -28.8016
18.8316 | -35.1286 | -28.8087
18.8377 | -35.1291 | -28.8157
18.8439 | -35.1296 | -28.8227
18.85 | -35.1301 | -28.8299
18.8561 | -35.1306 | -28.8371
18.8622 | -35.131 | -28.8441
18.8683 | -35.1315 | -28.8512
18.8744 | -35.132 | -28.8582
18.8805 | -35.1325 | -28.8654
18.8866 | -35.133 | -28.8727
18.8927 | -35.1334 | -28.8797
18.8988 | -35.1339 | -28.8867
18.9049 | -35.1344 | -28.8937
18.911 | -35.1348 | -28.9008
18.9171 | -35.1353 | -28.9079
18.9232 | -35.1358 | -28.9148
18.9293 | -35.1363 | -28.9219
18.9354 | -35.1368 | -28.929
18.9415 | -35.1373 | -28.936
18.9476 | -35.1377 | -28.943
18.9537 | -35.1382 | -28.9501
18.9598 | -35.1386 | -28.9572
18.9659 | -35.1391 | -28.9642
18.972 | -35.1396 | -28.9711
18.9781 | -35.1401 | -28.9783
18.9843 | -35.1405 | -28.9854
18.9904 | -35.141 | -28.9925
18.9965 | -35.1415 | -28.9994
19.0026 | -35.142 | -29.0065
19.0087 | -35.1425 | -29.0136
19.0148 | -35.1429 | -29.0207
19.0209 | -35.1434 | -29.0279
19.027 | -35.1438 | -29.035
19.0331 | -35.1443 | -29.042
19.0392 | -35.1448 | -29.0492
19.0453 | -35.1452 | -29.0563
19.0514 | -35.1457 | -29.0632
19.0575 | -35.1462 | -29.0703
19.0636 | -35.1466 | -29.0773
19.0697 | -35.1471 | -29.0844
19.0758 | -35.1476 | -29.0914
19.0819 | -35.148 | -29.0987
19.088 | -35.1485 | -29.1057
19.0941 | -35.149 | -29.1128
19.1002 | -35.1495 | -29.1198
19.1063 | -35.1499 | -29.1269
19.1124 | -35.1504 | -29.134
19.1185 | -35.1509 | -29.141
19.1246 | -35.1514 | -29.148
19.1308 | -35.1519 | -29.1553
19.1369 | -35.1523 | -29.1622
19.143 | -35.1528 | -29.1693
19.1491 | -35.1533 | -29.1764
19.1552 | -35.1538 | -29.1836
19.1613 | -35.1542 | -29.1907
19.1674 | -35.1547 | -29.1979
19.1735 | -35.1552 | -29.205
19.1796 | -35.1556 | -29.2121
19.1857 | -35.1561 | -29.2193
19.1918 | -35.1566 | -29.2263
19.1979 | -35.157 | -29.2335
19.204 | -35.1575 | -29.2406
19.2101 | -35.158 | -29.2475
19.2162 | -35.1585 | -29.2546
19.2223 | -35.1589 | -29.2616
19.2284 | -35.1594 | -29.2687
19.2345 | -35.1599 | -29.2757
19.2406 | -35.1604 | -29.2824
19.2467 | -35.1609 | -29.2895
19.2528 | -35.1613 | -29.2966
19.2589 | -35.1618 | -29.3037
19.265 | -35.1623 | -29.311
19.2712 | -35.1628 | -29.3181
19.2773 | -35.1633 | -29.325
19.2834 | -35.1637 | -29.332
19.2895 | -35.1642 | -29.3392
19.2956 | -35.1647 | -29.3464
19.3017 | -35.1652 | -29.3535
19.3078 | -35.1656 | -29.3607
19.3139 | -35.1661 | -29.3671
19.32 | -35.1666 | -29.374
19.3261 | -35.167 | -29.381
19.3322 | -35.1675 | -29.3879
19.3383 | -35.168 | -29.3949
19.3444 | -35.1684 | -29.4017
19.3505 | -35.1688 | -29.4087
19.3566 | -35.1693 | -29.4157
19.3627 | -35.1698 | -29.4228
19.3688 | -35.1703 | -29.4297
19.3749 | -35.1707 | -29.437
19.381 | -35.1712 | -29.4437
19.3871 | -35.1717 | -29.4507
19.3932 | -35.1721 | -29.4579
19.3993 | -35.1725 | -29.4651
19.4054 | -35.173 | -29.4722
19.4115 | -35.1735 | -29.4793
19.4177 | -35.1739 | -29.486
19.4238 | -35.1744 | -29.493
19.4299 | -35.1749 | -29.5
19.436 | -35.1753 | -29.5072
19.4421 | -35.1758 | -29.5141
19.4482 | -35.1763 | -29.5209
19.4543 | -35.1767 | -29.5276
19.4604 | -35.1772 | -29.5347
19.4665 | -35.1776 | -29.5417
19.4726 | -35.1781 | -29.5486
19.4787 | -35.1785 | -29.5557
19.4848 | -35.179 | -29.5628
19.4909 | -35.1795 | -29.57
19.497 | -35.1799 | -29.577
19.5031 | -35.1804 | -29.5839
19.5092 | -35.1808 | -29.5908
19.5153 | -35.1813 | -29.5976
19.5214 | -35.1817 | -29.6046
19.5275 | -35.1822 | -29.6116
19.5336 | -35.1826 | -29.6185
19.5397 | -35.1831 | -29.6255
19.5458 | -35.1835 | -29.6323
19.5519 | -35.184 | -29.6392
19.5581 | -35.1844 | -29.6462
19.5642 | -35.1849 | -29.6531
19.5703 | -35.1854 | -29.6601
19.5764 | -35.1859 | -29.6672
19.5825 | -35.1864 | -29.6741
19.5886 | -35.1868 | -29.6813
19.5947 | -35.1873 | -29.688
19.6008 | -35.1878 | -29.6949
19.6069 | -35.1882 | -29.7018
19.613 | -35.1887 | -29.7089
19.6191 | -35.1892 | -29.7157
19.6252 | -35.1896 | -29.7227
19.6313 | -35.1901 | -29.7298
19.6374 | -35.1905 | -29.7369
19.6435 | -35.191 | -29.7436
19.6496 | -35.1914 | -29.7505
19.6557 | -35.1919 | -29.7575
19.6618 | -35.1924 | -29.7639
19.6679 | -35.1929 | -29.7711
19.674 | -35.1933 | -29.7779
19.6801 | -35.1938 | -29.7849
19.6862 | -35.1943 | -29.7917
19.6923 | -35.1947 | -29.7983
19.6984 | -35.1952 | -29.8052
19.7046 | -35.1957 | -29.8123
19.7107 | -35.1961 | -29.8192
19.7168 | -35.1966 | -29.826
19.7229 | -35.1971 | -29.8328
19.729 | -35.1975 | -29.8396
19.7351 | -35.198 | -29.8462
19.7412 | -35.1985 | -29.8532
19.7473 | -35.199 | -29.8603
19.7534 | -35.1994 | -29.8675
19.7595 | -35.1999 | -29.8745
19.7656 | -35.2003 | -29.8813
19.7717 | -35.2008 | -29.888
19.7778 | -35.2012 | -29.895
19.7839 | -35.2017 | -29.9019
19.79 | -35.2022 | -29.9089
19.7961 | -35.2026 | -29.9158
19.8022 | -35.2031 | -29.9224
19.8083 | -35.2035 | -29.9294
19.8144 | -35.204 | -29.9362
19.8205 | -35.2045 | -29.9433
19.8266 | -35.205 | -29.9501
19.8327 | -35.2054 | -29.9569
19.8388 | -35.2059 | -29.9637
19.845 | -35.2064 | -29.9707
19.8511 | -35.2068 | -29.9778
19.8572 | -35.2073 | -29.9847
19.8633 | -35.2077 | -29.9917
19.8694 | -35.2082 | -29.9985
19.8755 | -35.2087 | -30.0051
19.8816 | -35.2092 | -30.0121
19.8877 | -35.2096 | -30.019
19.8938 | -35.2101 | -30.0259
19.8999 | -35.2106 | -30.0326
19.906 | -35.211 | -30.0395
19.9121 | -35.2114 | -30.0463
19.9182 | -35.2119 | -30.0533
19.9243 | -35.2123 | -30.0602
19.9304 | -35.2127 | -30.067
19.9365 | -35.2132 | -30.0739
19.9426 | -35.2137 | -30.0807
19.9487 | -35.2141 | -30.0876
19.9548 | -35.2145 | -30.0945
19.9609 | -35.215 | -30.1014
19.967 | -35.2154 | -30.1084
19.9731 | -35.2159 | -30.1151
19.9792 | -35.2164 | -30.1225
19.9853 | -35.2169 | -30.1288
19.9915 | -35.2173 | -30.1358
19.9976 | -35.2178 | -30.1429
20.0037 | -35.2182 | -30.1497
20.0098 | -35.2186 | -30.1566
20.0159 | -35.2191 | -30.1635
20.022 | -35.2196 | -30.1704
20.0281 | -35.22 | -30.1772
20.0342 | -35.2205 | -30.1844
20.0403 | -35.221 | -30.1913
20.0464 | -35.2214 | -30.1984
20.0525 | -35.2218 | -30.2055
20.0586 | -35.2223 | -30.2124
20.0647 | -35.2228 | -30.219
20.0708 | -35.2233 | -30.226
20.0769 | -35.2237 | -30.2331
20.083 | -35.2242 | -30.2395
20.0891 | -35.2246 | -30.2466
20.0952 | -35.2251 | -30.2532
20.1013 | -35.2255 | -30.2602
20.1074 | -35.226 | -30.267
20.1135 | -35.2265 | -30.2738
20.1196 | -35.2269 | -30.2802
20.1257 | -35.2274 | -30.2871
20.1319 | -35.2278 | -30.2939
20.138 | -35.2283 | -30.301
20.1441 | -35.2287 | -30.3078
20.1502 | -35.2292 | -30.3149
20.1563 | -35.2295 | -30.3213
20.1624 | -35.23 | -30.3282
20.1685 | -35.2304 | -30.3352
20.1746 | -35.2308 | -30.3418
20.1807 | -35.2313 | -30.3487
20.1868 | -35.2318 | -30.3558
20.1929 | -35.2321 | -30.3623
20.199 | -35.2326 | -30.369
20.2051 | -35.233 | -30.3759
20.2112 | -35.2335 | -30.3828
20.2173 | -35.2339 | -30.3898
20.2234 | -35.2343 | -30.3965
20.2295 | -35.2348 | -30.4034
20.2356 | -35.2352 | -30.4102
20.2417 | -35.2357 | -30.417
20.2478 | -35.2361 | -30.4239
20.2539 | -35.2366 | -30.4308
20.26 | -35.237 | -30.4374
20.2661 | -35.2375 | -30.4443
20.2723 | -35.2379 | -30.4511
20.2784 | -35.2384 | -30.4581
20.2845 | -35.2388 | -30.465
20.2906 | -35.2393 | -30.4716
20.2967 | -35.2397 | -30.4784
20.3028 | -35.2401 | -30.4852
20.3089 | -35.2406 | -30.492
20.315 | -35.241 | -30.4986
20.3211 | -35.2415 | -30.5054
20.3272 | -35.242 | -30.5123
20.3333 | -35.2425 | -30.5187
20.3394 | -35.243 | -30.5256
20.3455 | -35.2435 | -30.5325
20.3516 | -35.2439 | -30.5392
20.3577 | -35.2443 | -30.5461
20.3638 | -35.2448 | -30.5528
20.3699 | -35.2452 | -30.5595
20.376 | -35.2457 | -30.5661
20.3821 | -35.2462 | -30.5729
20.3882 | -35.2466 | -30.5796
20.3943 | -35.2471 | -30.5865
20.4004 | -35.2475 | -30.5929
20.4065 | -35.2479 | -30.5996
20.4126 | -35.2483 | -30.6065
20.4188 | -35.2487 | -30.6133
20.4249 | -35.2492 | -30.62
20.431 | -35.2496 | -30.6267
20.4371 | -35.25 | -30.6331
20.4432 | -35.2505 | -30.6399
20.4493 | -35.2509 | -30.6469
20.4554 | -35.2514 | -30.6534
20.4615 | -35.2518 | -30.6604
20.4676 | -35.2523 | -30.667
20.4737 | -35.2527 | -30.6738
20.4798 | -35.2531 | -30.6805
20.4859 | -35.2536 | -30.6874
20.492 | -35.2541 | -30.6944
20.4981 | -35.2545 | -30.7011
20.5042 | -35.255 | -30.7077
20.5103 | -35.2554 | -30.7142
20.5164 | -35.2559 | -30.721
20.5225 | -35.2563 | -30.7277
20.5286 | -35.2568 | -30.7344
20.5347 | -35.2572 | -30.7413
20.5408 | -35.2578 | -30.748
20.5469 | -35.2582 | -30.7549
20.553 | -35.2587 | -30.7617
20.5592 | -35.2592 | -30.7687
20.5653 | -35.2596 | -30.7754
20.5714 | -35.26 | -30.782
20.5775 | -35.2606 | -30.7884
20.5836 | -35.2611 | -30.7953
20.5897 | -35.2615 | -30.802
20.5958 | -35.2619 | -30.809
20.6019 | -35.2624 | -30.8157
20.608 | -35.2628 | -30.8227
20.6141 | -35.2632 | -30.829
20.6202 | -35.2636 | -30.8357
20.6263 | -35.264 | -30.8423
20.6324 | -35.2645 | -30.8492
20.6385 | -35.2649 | -30.856
20.6446 | -35.2654 | -30.8627
20.6507 | -35.2658 | -30.8693
20.6568 | -35.2663 | -30.876
20.6629 | -35.2667 | -30.8828
20.669 | -35.2671 | -30.8902
20.6751 | -35.2676 | -30.8971
20.6812 | -35.2681 | -30.9035
20.6873 | -35.2685 | -30.9104
20.6934 | -35.269 | -30.9171
20.6995 | -35.2694 | -30.9233
20.7057 | -35.2699 | -30.9295
20.7118 | -35.2704 | -30.9363
20.7179 | -35.2708 | -30.9429
20.724 | -35.2712 | -30.95
20.7301 | -35.2717 | -30.9566
20.7362 | -35.2721 | -30.9636
20.7423 | -35.2725 | -30.9703
20.7484 | -35.273 | -30.9768
20.7545 | -35.2735 | -30.9832
20.7606 | -35.2739 | -30.9902
20.7667 | -35.2744 | -30.9972
20.7728 | -35.2749 | -31.0037
20.7789 | -35.2753 | -31.0106
20.785 | -35.2758 | -31.0178
20.7911 | -35.2762 | -31.0243
20.7972 | -35.2767 | -31.0314
20.8033 | -35.2771 | -31.0382
20.8094 | -35.2776 | -31.0448
20.8155 | -35.278 | -31.0516
20.8216 | -35.2785 | -31.0581
20.8277 | -35.2789 | -31.0649
20.8338 | -35.2794 | -31.0715
20.8399 | -35.2798 | -31.0778
20.8461 | -35.2803 | -31.0849
20.8522 | -35.2807 | -31.0912
20.8583 | -35.2812 | -31.0979
20.8644 | -35.2816 | -31.1047
20.8705 | -35.282 | -31.1114
20.8766 | -35.2825 | -31.118
20.8827 | -35.2829 | -31.1246
20.8888 | -35.2834 | -31.1312
20.8949 | -35.2838 | -31.1378
20.901 | -35.2842 | -31.1444
20.9071 | -35.2847 | -31.1509
20.9132 | -35.2851 | -31.1578
20.9193 | -35.2856 | -31.1643
20.9254 | -35.2861 | -31.171
20.9315 | -35.2865 | -31.1776
20.9376 | -35.287 | -31.1843
20.9437 | -35.2874 | -31.1906
20.9498 | -35.2879 | -31.197
20.9559 | -35.2883 | -31.2034
20.962 | -35.2888 | -31.21
20.9681 | -35.2892 | -31.2167
20.9742 | -35.2895 | -31.2233
20.9803 | -35.29 | -31.2299
20.9864 | -35.2904 | -31.2363
20.9926 | -35.2909 | -31.2426
20.9987 | -35.2914 | -31.2488
21.0048 | -35.2919 | -31.2548
21.0109 | -35.2924 | -31.2616
21.017 | -35.2928 | -31.2688
21.0231 | -35.2933 | -31.276
21.0292 | -35.2938 | -31.2831
21.0353 | -35.2942 | -31.2895
21.0414 | -35.2947 | -31.2957
21.0475 | -35.2951 | -31.302
21.0536 | -35.2955 | -31.3083
21.0597 | -35.296 | -31.3148
21.0658 | -35.2964 | -31.3212
21.0719 | -35.2968 | -31.3277
21.078 | -35.2972 | -31.3347
21.0841 | -35.2977 | -31.3415
21.0902 | -35.2981 | -31.3479
21.0963 | -35.2985 | -31.3542
21.1024 | -35.2988 | -31.3608
21.1085 | -35.2992 | -31.3681
21.1146 | -35.2996 | -31.3747
21.1207 | -35.3001 | -31.3811
21.1268 | -35.3005 | -31.3872
21.133 | -35.301 | -31.3938
21.1391 | -35.3014 | -31.4001
21.1452 | -35.3018 | -31.4062
21.1513 | -35.3023 | -31.4131
21.1574 | -35.3027 | -31.4197
21.1635 | -35.3031 | -31.4264
21.1696 | -35.3036 | -31.4331
21.1757 | -35.304 | -31.4392
21.1818 | -35.3044 | -31.4461
21.1879 | -35.3049 | -31.4528
21.194 | -35.3053 | -31.4594
21.2001 | -35.3057 | -31.4657
21.2062 | -35.3061 | -31.4726
21.2123 | -35.3066 | -31.4789
21.2184 | -35.307 | -31.4854
21.2245 | -35.3075 | -31.4923
21.2306 | -35.308 | -31.4988
21.2367 | -35.3084 | -31.5049
21.2428 | -35.3088 | -31.5123
21.2489 | -35.3094 | -31.5189
21.255 | -35.3097 | -31.5259
21.2611 | -35.3102 | -31.5327
21.2672 | -35.3105 | -31.5394
21.2733 | -35.311 | -31.5465
21.2795 | -35.3114 | -31.5528
21.2856 | -35.3119 | -31.5593
21.2917 | -35.3124 | -31.5658
21.2978 | -35.3128 | -31.5725
21.3039 | -35.3132 | -31.5791
21.31 | -35.3137 | -31.5856
21.3161 | -35.314 | -31.5914
21.3222 | -35.3145 | -31.5979
21.3283 | -35.3149 | -31.6047
21.3344 | -35.3153 | -31.6115
21.3405 | -35.3158 | -31.6178
21.3466 | -35.3163 | -31.624
21.3527 | -35.3168 | -31.6299
21.3588 | -35.3172 | -31.636
21.3649 | -35.3176 | -31.6428
21.371 | -35.3181 | -31.649
21.3771 | -35.3186 | -31.6552
21.3832 | -35.319 | -31.6616
21.3893 | -35.3194 | -31.6685
21.3954 | -35.3199 | -31.6747
21.4015 | -35.3204 | -31.6815
21.4076 | -35.3208 | -31.6882
21.4137 | -35.3212 | -31.6947
21.4199 | -35.3217 | -31.7013
21.426 | -35.3221 | -31.7074
21.4321 | -35.3225 | -31.714
21.4382 | -35.3229 | -31.7205
21.4443 | -35.3233 | -31.7271
21.4504 | -35.3238 | -31.7337
21.4565 | -35.3242 | -31.74
21.4626 | -35.3247 | -31.7462
21.4687 | -35.3251 | -31.7526
21.4748 | -35.3256 | -31.7588
21.4809 | -35.326 | -31.7648
21.487 | -35.3265 | -31.7708
21.4931 | -35.327 | -31.7771
21.4992 | -35.3275 | -31.7835
21.5053 | -35.3278 | -31.79
21.5114 | -35.3283 | -31.7965
21.5175 | -35.3288 | -31.8031
21.5236 | -35.3292 | -31.8098
21.5297 | -35.3298 | -31.8166
21.5358 | -35.3302 | -31.8232
21.5419 | -35.3306 | -31.8297
21.548 | -35.3311 | -31.8362
21.5541 | -35.3315 | -31.8429
21.5602 | -35.332 | -31.8496
21.5664 | -35.3325 | -31.8555
21.5725 | -35.333 | -31.8621
21.5786 | -35.3335 | -31.8689
21.5847 | -35.3339 | -31.8757
21.5908 | -35.3344 | -31.882
21.5969 | -35.3348 | -31.8886
21.603 | -35.3352 | -31.8952
21.6091 | -35.3357 | -31.9021
21.6152 | -35.3361 | -31.909
21.6213 | -35.3365 | -31.9156
21.6274 | -35.337 | -31.9224
21.6335 | -35.3375 | -31.9284
21.6396 | -35.3379 | -31.9353
21.6457 | -35.3384 | -31.9424
21.6518 | -35.3388 | -31.9485
21.6579 | -35.3393 | -31.955
21.664 | -35.3397 | -31.9615
21.6701 | -35.34 | -31.9673
21.6762 | -35.3405 | -31.9741
21.6823 | -35.3409 | -31.9809
21.6884 | -35.3413 | -31.9877
21.6945 | -35.3418 | -31.9941
21.7006 | -35.3423 | -32.0004
21.7068 | -35.3427 | -32.0068
21.7129 | -35.3432 | -32.013
21.719 | -35.3436 | -32.0193
21.7251 | -35.344 | -32.0256
21.7312 | -35.3444 | -32.0321
21.7373 | -35.3449 | -32.0381
21.7434 | -35.3454 | -32.0444
21.7495 | -35.3458 | -32.0504
21.7556 | -35.3462 | -32.0565
21.7617 | -35.3467 | -32.0631
21.7678 | -35.3471 | -32.0695
21.7739 | -35.3476 | -32.075
21.78 | -35.348 | -32.0816
21.7861 | -35.3485 | -32.0881
21.7922 | -35.3489 | -32.095
21.7983 | -35.3493 | -32.1016
21.8044 | -35.3497 | -32.108
21.8105 | -35.3501 | -32.1146
21.8166 | -35.3506 | -32.121
21.8227 | -35.351 | -32.1275
21.8288 | -35.3515 | -32.1337
21.8349 | -35.3519 | -32.1401
21.841 | -35.3523 | -32.1465
21.8471 | -35.3528 | -32.1524
21.8533 | -35.3533 | -32.1589
21.8594 | -35.3538 | -32.1653
21.8655 | -35.3543 | -32.1716
21.8716 | -35.3547 | -32.1778
21.8777 | -35.3552 | -32.184
21.8838 | -35.3556 | -32.1906
21.8899 | -35.356 | -32.1973
21.896 | -35.3565 | -32.2033
21.9021 | -35.3569 | -32.2093
21.9082 | -35.3573 | -32.216
21.9143 | -35.3577 | -32.2226
21.9204 | -35.3582 | -32.2293
21.9265 | -35.3586 | -32.2355
21.9326 | -35.3591 | -32.2423
21.9387 | -35.3595 | -32.2488
21.9448 | -35.36 | -32.2552
21.9509 | -35.3605 | -32.2616
21.957 | -35.3609 | -32.2679
21.9631 | -35.3613 | -32.2747
21.9692 | -35.3617 | -32.2809
21.9753 | -35.3622 | -32.287
21.9814 | -35.3626 | -32.2933
21.9875 | -35.363 | -32.2994
21.9937 | -35.3635 | -32.3057
21.9998 | -35.3639 | -32.3119
22.0059 | -35.3644 | -32.318
22.012 | -35.3648 | -32.3242
22.0181 | -35.3652 | -32.3305
22.0242 | -35.3656 | -32.3372
22.0303 | -35.3661 | -32.3434
22.0364 | -35.3665 | -32.3498
22.0425 | -35.367 | -32.3558
22.0486 | -35.3673 | -32.3622
22.0547 | -35.3678 | -32.3683
22.0608 | -35.3682 | -32.3744
22.0669 | -35.3686 | -32.3803
22.073 | -35.3691 | -32.3866
22.0791 | -35.3695 | -32.3932
22.0852 | -35.3699 | -32.3997
22.0913 | -35.3703 | -32.4056
22.0974 | -35.3707 | -32.4121
22.1035 | -35.3711 | -32.4186
22.1096 | -35.3716 | -32.4249
22.1157 | -35.372 | -32.4314
22.1218 | -35.3724 | -32.4379
22.1279 | -35.3728 | -32.4437
22.134 | -35.3732 | -32.45
22.1402 | -35.3736 | -32.4564
22.1463 | -35.374 | -32.4629
22.1524 | -35.3745 | -32.4695
22.1585 | -35.3749 | -32.4756
22.1646 | -35.3754 | -32.4814
22.1707 | -35.3758 | -32.4877
22.1768 | -35.3763 | -32.4939
22.1829 | -35.3767 | -32.5005
22.189 | -35.3772 | -32.507
22.1951 | -35.3776 | -32.5129
22.2012 | -35.378 | -32.5192
22.2073 | -35.3785 | -32.5254
22.2134 | -35.3789 | -32.5319
22.2195 | -35.3794 | -32.5381
22.2256 | -35.3799 | -32.5441
22.2317 | -35.3804 | -32.5507
22.2378 | -35.3808 | -32.5568
22.2439 | -35.3812 | -32.5632
22.25 | -35.3816 | -32.5694
22.2561 | -35.3821 | -32.5755
22.2622 | -35.3825 | -32.5819
22.2683 | -35.383 | -32.5876
22.2744 | -35.3835 | -32.5937
22.2806 | -35.3839 | -32.6006
22.2867 | -35.3843 | -32.6068
22.2928 | -35.3847 | -32.6128
22.2989 | -35.3852 | -32.6192
22.305 | -35.3857 | -32.6256
22.3111 | -35.3862 | -32.6319
22.3172 | -35.3867 | -32.6381
22.3233 | -35.3872 | -32.6444
22.3294 | -35.3875 | -32.6508
22.3355 | -35.388 | -32.6572
22.3416 | -35.3884 | -32.6636
22.3477 | -35.3889 | -32.67
22.3538 | -35.3893 | -32.6765
22.3599 | -35.3898 | -32.6832
22.366 | -35.3901 | -32.6896
22.3721 | -35.3906 | -32.6952
22.3782 | -35.391 | -32.7016
22.3843 | -35.3914 | -32.7078
22.3904 | -35.3919 | -32.7141
22.3965 | -35.3924 | -32.7205
22.4026 | -35.3928 | -32.727
22.4087 | -35.3932 | -32.733
22.4148 | -35.3937 | -32.7397
22.4209 | -35.3941 | -32.7458
22.4271 | -35.3945 | -32.7522
22.4332 | -35.3951 | -32.7587
22.4393 | -35.3955 | -32.7652
22.4454 | -35.3959 | -32.7708
22.4515 | -35.3963 | -32.777
22.4576 | -35.3967 | -32.7835
22.4637 | -35.3971 | -32.79
22.4698 | -35.3976 | -32.7961
22.4759 | -35.398 | -32.8018
22.482 | -35.3984 | -32.808
22.4881 | -35.3988 | -32.8141
22.4942 | -35.3992 | -32.8206
22.5003 | -35.3996 | -32.8268
22.5064 | -35.4001 | -32.8328
22.5125 | -35.4005 | -32.839
22.5186 | -35.4009 | -32.8454
22.5247 | -35.4014 | -32.8517
22.5308 | -35.4018 | -32.8583
22.5369 | -35.4022 | -32.8647
22.543 | -35.4027 | -32.8708
22.5491 | -35.4032 | -32.8768
22.5552 | -35.4036 | -32.8832
22.5613 | -35.404 | -32.8894
22.5675 | -35.4045 | -32.8955
22.5736 | -35.4049 | -32.9021
22.5797 | -35.4053 | -32.9082
22.5858 | -35.4058 | -32.9141
22.5919 | -35.4063 | -32.9203
22.598 | -35.4067 | -32.9265
22.6041 | -35.4071 | -32.9328
22.6102 | -35.4076 | -32.9394
22.6163 | -35.408 | -32.9458
22.6224 | -35.4085 | -32.9522
22.6285 | -35.409 | -32.9588
22.6346 | -35.4094 | -32.965
22.6407 | -35.4099 | -32.9711
22.6468 | -35.4103 | -32.9776
22.6529 | -35.4108 | -32.9835
22.659 | -35.4112 | -32.9894
22.6651 | -35.4115 | -32.9958
22.6712 | -35.412 | -33.0021
22.6773 | -35.4125 | -33.0085
22.6834 | -35.4129 | -33.0149
22.6895 | -35.4133 | -33.0207
22.6956 | -35.4139 | -33.0269
22.7017 | -35.4142 | -33.0331
22.7079 | -35.4147 | -33.039
22.714 | -35.4151 | -33.0452
22.7201 | -35.4155 | -33.0514
22.7262 | -35.416 | -33.0576
22.7323 | -35.4164 | -33.0637
22.7384 | -35.4169 | -33.0701
22.7445 | -35.4172 | -33.0765
22.7506 | -35.4176 | -33.0829
22.7567 | -35.418 | -33.0891
22.7628 | -35.4184 | -33.0955
22.7689 | -35.4189 | -33.1015
22.775 | -35.4193 | -33.1079
22.7811 | -35.4197 | -33.1141
22.7872 | -35.4201 | -33.1207
22.7933 | -35.4206 | -33.1263
22.7994 | -35.421 | -33.1326
22.8055 | -35.4215 | -33.139
22.8116 | -35.422 | -33.1455
22.8177 | -35.4225 | -33.1519
22.8238 | -35.4229 | -33.1583
22.8299 | -35.4234 | -33.1643
22.836 | -35.4238 | -33.1702
22.8421 | -35.4242 | -33.1768
22.8482 | -35.4246 | -33.1832
22.8544 | -35.425 | -33.1893
22.8605 | -35.4255 | -33.1956
22.8666 | -35.4259 | -33.2019
22.8727 | -35.4263 | -33.2081
22.8788 | -35.4268 | -33.2145
22.8849 | -35.4272 | -33.2208
22.891 | -35.4276 | -33.2269
22.8971 | -35.4281 | -33.2327
22.9032 | -35.4285 | -33.2387
22.9093 | -35.429 | -33.2447
22.9154 | -35.4294 | -33.2506
22.9215 | -35.4298 | -33.2566
22.9276 | -35.4302 | -33.2626
22.9337 | -35.4306 | -33.2686
22.9398 | -35.4311 | -33.275
22.9459 | -35.4315 | -33.2813
22.952 | -35.4319 | -33.2872
22.9581 | -35.4323 | -33.2934
22.9642 | -35.4328 | -33.2996
22.9703 | -35.4332 | -33.3055
22.9764 | -35.4335 | -33.3118
22.9825 | -35.4339 | -33.318
22.9886 | -35.4344 | -33.3242
22.9948 | -35.4348 | -33.3305
23.0009 | -35.4352 | -33.3369
23.007 | -35.4356 | -33.343
23.0131 | -35.436 | -33.3493
23.0192 | -35.4365 | -33.3556
23.0253 | -35.4369 | -33.3617
23.0314 | -35.4374 | -33.3678
23.0375 | -35.4378 | -33.374
23.0436 | -35.4381 | -33.3803
23.0497 | -35.4386 | -33.3865
23.0558 | -35.4391 | -33.3925
23.0619 | -35.4395 | -33.3985
23.068 | -35.4399 | -33.4044
23.0741 | -35.4403 | -33.4105
23.0802 | -35.4408 | -33.4167
23.0863 | -35.4412 | -33.4232
23.0924 | -35.4416 | -33.4294
23.0985 | -35.442 | -33.4358
23.1046 | -35.4424 | -33.4422
23.1107 | -35.4428 | -33.4483
23.1168 | -35.4432 | -33.4541
23.1229 | -35.4437 | -33.4603
23.129 | -35.4441 | -33.4665
23.1351 | -35.4445 | -33.4726
23.1413 | -35.4449 | -33.4789
23.1474 | -35.4453 | -33.4848
23.1535 | -35.4458 | -33.4909
23.1596 | -35.4462 | -33.4969
23.1657 | -35.4466 | -33.5029
23.1718 | -35.4471 | -33.509
23.1779 | -35.4476 | -33.5152
23.184 | -35.4481 | -33.5211
23.1901 | -35.4485 | -33.5271
23.1962 | -35.4489 | -33.5331
23.2023 | -35.4494 | -33.5396
23.2084 | -35.4498 | -33.5458
23.2145 | -35.4503 | -33.551
23.2206 | -35.4507 | -33.5571
23.2267 | -35.4511 | -33.5633
23.2328 | -35.4515 | -33.5695
23.2389 | -35.4519 | -33.5758
23.245 | -35.4523 | -33.5816
23.2511 | -35.4528 | -33.5871
23.2572 | -35.4532 | -33.5931
23.2633 | -35.4536 | -33.5992
23.2694 | -35.4541 | -33.6057
23.2755 | -35.4545 | -33.6118
23.2817 | -35.4549 | -33.6177
23.2878 | -35.4554 | -33.6237
23.2939 | -35.4558 | -33.6302
23.3 | -35.4563 | -33.6365
23.3061 | -35.4567 | -33.6429
23.3122 | -35.4572 | -33.649
23.3183 | -35.4575 | -33.6551
23.3244 | -35.458 | -33.6613
23.3305 | -35.4584 | -33.6673
23.3366 | -35.4589 | -33.6736
23.3427 | -35.4593 | -33.6799
23.3488 | -35.4597 | -33.6856
23.3549 | -35.4602 | -33.6918
23.361 | -35.4606 | -33.6978
23.3671 | -35.461 | -33.7041
23.3732 | -35.4614 | -33.71
23.3793 | -35.4618 | -33.7162
23.3854 | -35.4623 | -33.7226
23.3915 | -35.4627 | -33.7286
23.3976 | -35.4632 | -33.7349
23.4037 | -35.4636 | -33.7411
23.4098 | -35.464 | -33.7469
23.4159 | -35.4645 | -33.7529
23.422 | -35.4649 | -33.7591
23.4282 | -35.4653 | -33.7646
23.4343 | -35.4658 | -33.7706
23.4404 | -35.4662 | -33.7767
23.4465 | -35.4667 | -33.7832
23.4526 | -35.4672 | -33.7894
23.4587 | -35.4676 | -33.7955
23.4648 | -35.4679 | -33.8016
23.4709 | -35.4683 | -33.8077
23.477 | -35.4687 | -33.8137
23.4831 | -35.4691 | -33.8195
23.4892 | -35.4695 | -33.8258
23.4953 | -35.47 | -33.832
23.5014 | -35.4704 | -33.8381
23.5075 | -35.4708 | -33.8441
23.5136 | -35.4712 | -33.8501
23.5197 | -35.4717 | -33.8561
23.5258 | -35.4721 | -33.8624
23.5319 | -35.4725 | -33.8681
23.538 | -35.4729 | -33.8742
23.5441 | -35.4733 | -33.8802
23.5502 | -35.4737 | -33.8864
23.5563 | -35.4741 | -33.8925
23.5624 | -35.4746 | -33.8987
23.5686 | -35.475 | -33.9047
23.5747 | -35.4755 | -33.9102
23.5808 | -35.4759 | -33.9162
23.5869 | -35.4763 | -33.9222
23.593 | -35.4767 | -33.9283
23.5991 | -35.4771 | -33.9348
23.6052 | -35.4775 | -33.9406
23.6113 | -35.478 | -33.9467
23.6174 | -35.4784 | -33.9527
23.6235 | -35.4788 | -33.959
23.6296 | -35.4792 | -33.9651
23.6357 | -35.4796 | -33.9715
23.6418 | -35.4801 | -33.9773
23.6479 | -35.4805 | -33.9832
23.654 | -35.4809 | -33.9891
23.6601 | -35.4813 | -33.9947
23.6662 | -35.4817 | -34.0007
23.6723 | -35.4822 | -34.0067
23.6784 | -35.4826 | -34.0129
23.6845 | -35.483 | -34.0189
23.6906 | -35.4835 | -34.0249
23.6967 | -35.484 | -34.0307
23.7028 | -35.4844 | -34.0367
23.7089 | -35.4848 | -34.0423
23.7151 | -35.4852 | -34.0485
23.7212 | -35.4857 | -34.0546
23.7273 | -35.4861 | -34.0609
23.7334 | -35.4865 | -34.0672
23.7395 | -35.4869 | -34.0734
23.7456 | -35.4873 | -34.0798
23.7517 | -35.4877 | -34.0859
23.7578 | -35.4881 | -34.0918
23.7639 | -35.4886 | -34.0979
23.77 | -35.489 | -34.1038
23.7761 | -35.4895 | -34.11
23.7822 | -35.4898 | -34.1159
23.7883 | -35.4903 | -34.1219
23.7944 | -35.4907 | -34.1282
23.8005 | -35.4911 | -34.1342
23.8066 | -35.4915 | -34.1403
23.8127 | -35.492 | -34.1462
23.8188 | -35.4924 | -34.1524
23.8249 | -35.4929 | -34.1588
23.831 | -35.4933 | -34.1644
23.8371 | -35.4937 | -34.1707
23.8432 | -35.4941 | -34.1765
23.8493 | -35.4946 | -34.1823
23.8555 | -35.495 | -34.1882
23.8616 | -35.4954 | -34.1944
23.8677 | -35.4958 | -34.2005
23.8738 | -35.4962 | -34.2064
23.8799 | -35.4967 | -34.2126
23.886 | -35.497 | -34.2186
23.8921 | -35.4975 | -34.2245
23.8982 | -35.4979 | -34.2301
23.9043 | -35.4983 | -34.2365
23.9104 | -35.4987 | -34.242
23.9165 | -35.4991 | -34.2482
23.9226 | -35.4995 | -34.2543
23.9287 | -35.4999 | -34.2602
23.9348 | -35.5003 | -34.2666
23.9409 | -35.5007 | -34.2725
23.947 | -35.5011 | -34.279
23.9531 | -35.5014 | -34.2848
23.9592 | -35.5019 | -34.2908
23.9653 | -35.5023 | -34.2968
23.9714 | -35.5027 | -34.3025
23.9775 | -35.5031 | -34.3084
23.9836 | -35.5035 | -34.3142
23.9897 | -35.5037 | -34.32
23.9958 | -35.5041 | -34.326
24.002 | -35.5045 | -34.332
24.0081 | -35.5049 | -34.3381
24.0142 | -35.5054 | -34.344
24.0203 | -35.5058 | -34.3499
24.0264 | -35.5062 | -34.3559
24.0325 | -35.5066 | -34.362
24.0386 | -35.507 | -34.3679
24.0447 | -35.5073 | -34.3737
24.0508 | -35.5078 | -34.3795
24.0569 | -35.5082 | -34.3856
24.063 | -35.5086 | -34.3915
24.0691 | -35.509 | -34.3974
24.0752 | -35.5094 | -34.4029
24.0813 | -35.5098 | -34.4091
24.0874 | -35.5102 | -34.4153
24.0935 | -35.5107 | -34.4218
24.0996 | -35.511 | -34.4276
24.1057 | -35.5114 | -34.4333
24.1118 | -35.5118 | -34.4391
24.1179 | -35.5122 | -34.4454
24.124 | -35.5126 | -34.4513
24.1301 | -35.5131 | -34.4567
24.1362 | -35.5135 | -34.4627
24.1424 | -35.5139 | -34.4686
24.1485 | -35.5143 | -34.4746
24.1546 | -35.5147 | -34.4806
24.1607 | -35.515 | -34.4868
24.1668 | -35.5154 | -34.492
24.1729 | -35.5158 | -34.4981
24.179 | -35.5162 | -34.504
24.1851 | -35.5166 | -34.5102
24.1912 | -35.517 | -34.5163
24.1973 | -35.5173 | -34.5222
24.2034 | -35.5177 | -34.5281
24.2095 | -35.5182 | -34.5338
24.2156 | -35.5185 | -34.5397
24.2217 | -35.5189 | -34.5454
24.2278 | -35.5194 | -34.5512
24.2339 | -35.5198 | -34.5571
24.24 | -35.5201 | -34.5627
24.2461 | -35.5206 | -34.5686
24.2522 | -35.5209 | -34.5746
24.2583 | -35.5214 | -34.5803
24.2644 | -35.5218 | -34.5864
24.2705 | -35.5221 | -34.5926
24.2766 | -35.5225 | -34.5984
24.2827 | -35.5229 | -34.6045
24.2889 | -35.5234 | -34.6105
24.295 | -35.5237 | -34.6165
24.3011 | -35.5242 | -34.6222
24.3072 | -35.5246 | -34.628
24.3133 | -35.525 | -34.6338
24.3194 | -35.5253 | -34.6399
24.3255 | -35.5257 | -34.6458
24.3316 | -35.5261 | -34.652
24.3377 | -35.5265 | -34.6577
24.3438 | -35.5269 | -34.663
24.3499 | -35.5273 | -34.6688
24.356 | -35.5277 | -34.6748
24.3621 | -35.5281 | -34.6804
24.3682 | -35.5285 | -34.6864
24.3743 | -35.5289 | -34.6921
24.3804 | -35.5293 | -34.6979
24.3865 | -35.5297 | -34.7038
24.3926 | -35.5301 | -34.7098
24.3987 | -35.5304 | -34.7157
24.4048 | -35.5309 | -34.7217
24.4109 | -35.5312 | -34.7272
24.417 | -35.5316 | -34.7333
24.4231 | -35.532 | -34.7387
24.4293 | -35.5324 | -34.7449
24.4354 | -35.5328 | -34.7507
24.4415 | -35.5333 | -34.7564
24.4476 | -35.5337 | -34.7618
24.4537 | -35.5341 | -34.7675
24.4598 | -35.5345 | -34.7731
24.4659 | -35.5349 | -34.7789
24.472 | -35.5353 | -34.7849
24.4781 | -35.5357 | -34.7908
24.4842 | -35.536 | -34.7967
24.4903 | -35.5364 | -34.8026
24.4964 | -35.5368 | -34.8084
24.5025 | -35.5371 | -34.8142
24.5086 | -35.5375 | -34.8201
24.5147 | -35.5379 | -34.826
24.5208 | -35.5384 | -34.8319
24.5269 | -35.5387 | -34.8375
24.533 | -35.5392 | -34.8435
24.5391 | -35.5395 | -34.8495
24.5452 | -35.5399 | -34.8554
24.5513 | -35.5403 | -34.8611
24.5574 | -35.5407 | -34.8669
24.5635 | -35.541 | -34.8726
24.5696 | -35.5414 | -34.8782
24.5758 | -35.5418 | -34.8843
24.5819 | -35.5422 | -34.8901
24.588 | -35.5425 | -34.8957
24.5941 | -35.5429 | -34.9016
24.6002 | -35.5432 | -34.9076
24.6063 | -35.5437 | -34.9138
24.6124 | -35.5441 | -34.9199
24.6185 | -35.5445 | -34.9258
24.6246 | -35.5448 | -34.9309
24.6307 | -35.5452 | -34.9369
24.6368 | -35.5456 | -34.9426
24.6429 | -35.546 | -34.9486
24.649 | -35.5464 | -34.9543
24.6551 | -35.5467 | -34.9602
24.6612 | -35.5471 | -34.9655
24.6673 | -35.5475 | -34.9714
24.6734 | -35.5479 | -34.9772
24.6795 | -35.5483 | -34.9832
24.6856 | -35.5487 | -34.9889
24.6917 | -35.5491 | -34.9948
24.6978 | -35.5495 | -35.0004
24.7039 | -35.5499 | -35.0063
24.71 | -35.5503 | -35.0123
24.7162 | -35.5506 | -35.0181
24.7223 | -35.551 | -35.0239
24.7284 | -35.5514 | -35.0296
24.7345 | -35.5518 | -35.0354
24.7406 | -35.5522 | -35.0411
24.7467 | -35.5526 | -35.0469
24.7528 | -35.5529 | -35.0526
24.7589 | -35.5533 | -35.0583
24.765 | -35.5537 | -35.0636
24.7711 | -35.5541 | -35.0692
24.7772 | -35.5544 | -35.0752
24.7833 | -35.5548 | -35.0808
24.7894 | -35.5552 | -35.0865
24.7955 | -35.5556 | -35.0922
24.8016 | -35.5559 | -35.0978
24.8077 | -35.5563 | -35.1038
24.8138 | -35.5567 | -35.1096
24.8199 | -35.5571 | -35.1154
24.826 | -35.5575 | -35.1211
24.8321 | -35.5578 | -35.1267
24.8382 | -35.5582 | -35.132
24.8443 | -35.5586 | -35.1379
24.8504 | -35.559 | -35.1438
24.8565 | -35.5594 | -35.1496
24.8627 | -35.5598 | -35.1551
24.8688 | -35.5601 | -35.1605
24.8749 | -35.5605 | -35.1662
24.881 | -35.5608 | -35.1722
24.8871 | -35.5612 | -35.1782
24.8932 | -35.5616 | -35.184
24.8993 | -35.5619 | -35.1896
24.9054 | -35.5622 | -35.1953
24.9115 | -35.5626 | -35.2009
24.9176 | -35.563 | -35.2067
24.9237 | -35.5634 | -35.2124
24.9298 | -35.5638 | -35.2184
24.9359 | -35.5642 | -35.2244
24.942 | -35.5645 | -35.23
24.9481 | -35.5649 | -35.2358
24.9542 | -35.5653 | -35.2414
24.9603 | -35.5657 | -35.247
24.9664 | -35.566 | -35.2524
24.9725 | -35.5664 | -35.2581
24.9786 | -35.5667 | -35.2638
24.9847 | -35.567 | -35.2692
24.9908 | -35.5674 | -35.2748
24.9969 | -35.5678 | -35.2809
25.0031 | -35.5682 | -35.2865
25.0092 | -35.5685 | -35.2919
25.0153 | -35.5689 | -35.2976
25.0214 | -35.5693 | -35.3033
25.0275 | -35.5697 | -35.3089
25.0336 | -35.5701 | -35.3147
25.0397 | -35.5704 | -35.3207
25.0458 | -35.5708 | -35.3262
25.0519 | -35.5712 | -35.3319
25.058 | -35.5715 | -35.3376
25.0641 | -35.5719 | -35.343
25.0702 | -35.5723 | -35.3486
25.0763 | -35.5727 | -35.3542
25.0824 | -35.573 | -35.3597
25.0885 | -35.5733 | -35.3655
25.0946 | -35.5737 | -35.3714
25.1007 | -35.5741 | -35.3769
25.1068 | -35.5745 | -35.3825
25.1129 | -35.5749 | -35.3882
25.119 | -35.5752 | -35.3937
25.1251 | -35.5756 | -35.3993
25.1312 | -35.576 | -35.4048
25.1373 | -35.5764 | -35.4109
25.1435 | -35.5767 | -35.4167
25.1496 | -35.5771 | -35.4219
25.1557 | -35.5774 | -35.4276
25.1618 | -35.5777 | -35.4332
25.1679 | -35.5781 | -35.439
25.174 | -35.5784 | -35.4445
25.1801 | -35.5788 | -35.4505
25.1862 | -35.5791 | -35.4558
25.1923 | -35.5794 | -35.4619
25.1984 | -35.5798 | -35.4674
25.2045 | -35.5802 | -35.4731
25.2106 | -35.5806 | -35.4787
25.2167 | -35.5809 | -35.4844
25.2228 | -35.5812 | -35.49
25.2289 | -35.5816 | -35.4958
25.235 | -35.582 | -35.5017
25.2411 | -35.5824 | -35.5073
25.2472 | -35.5827 | -35.513
25.2533 | -35.5831 | -35.5188
25.2594 | -35.5834 | -35.5242
25.2655 | -35.5837 | -35.5297
25.2716 | -35.5841 | -35.5351
25.2777 | -35.5844 | -35.5407
25.2838 | -35.5848 | -35.5462
25.29 | -35.5852 | -35.552
25.2961 | -35.5855 | -35.5574
25.3022 | -35.5859 | -35.5629
25.3083 | -35.5863 | -35.5686
25.3144 | -35.5866 | -35.5742
25.3205 | -35.587 | -35.5797
25.3266 | -35.5873 | -35.5852
25.3327 | -35.5877 | -35.5909
25.3388 | -35.5881 | -35.5969
25.3449 | -35.5884 | -35.6025
25.351 | -35.5887 | -35.6083
25.3571 | -35.5891 | -35.614
25.3632 | -35.5894 | -35.6195
25.3693 | -35.5897 | -35.6252
25.3754 | -35.5901 | -35.6309
25.3815 | -35.5904 | -35.6365
25.3876 | -35.5907 | -35.642
25.3937 | -35.5911 | -35.6475
25.3998 | -35.5914 | -35.6526
25.4059 | -35.5917 | -35.6584
25.412 | -35.5921 | -35.6639
25.4181 | -35.5925 | -35.6694
25.4242 | -35.5928 | -35.6752
25.4304 | -35.5932 | -35.6809
25.4365 | -35.5936 | -35.6862
25.4426 | -35.5939 | -35.6919
25.4487 | -35.5943 | -35.6977
25.4548 | -35.5946 | -35.7034
25.4609 | -35.595 | -35.7091
25.467 | -35.5953 | -35.7145
25.4731 | -35.5957 | -35.7201
25.4792 | -35.596 | -35.7254
25.4853 | -35.5964 | -35.7308
25.4914 | -35.5968 | -35.7364
25.4975 | -35.5971 | -35.7419
25.5036 | -35.5974 | -35.7473
25.5097 | -35.5978 | -35.7529
25.5158 | -35.5981 | -35.7584
25.5219 | -35.5985 | -35.7642
25.528 | -35.5989 | -35.7698
25.5341 | -35.5992 | -35.7755
25.5402 | -35.5995 | -35.7809
25.5463 | -35.5999 | -35.7867
25.5524 | -35.6002 | -35.7921
25.5585 | -35.6006 | -35.7977
25.5646 | -35.6009 | -35.8034
25.5707 | -35.6013 | -35.8091
25.5769 | -35.6017 | -35.8146
25.583 | -35.602 | -35.82
25.5891 | -35.6023 | -35.8255
25.5952 | -35.6027 | -35.8312
25.6013 | -35.6031 | -35.8368
25.6074 | -35.6034 | -35.8422
25.6135 | -35.6037 | -35.8477
25.6196 | -35.6041 | -35.8534
25.6257 | -35.6045 | -35.8589
25.6318 | -35.6048 | -35.8644
25.6379 | -35.6051 | -35.8701
25.644 | -35.6055 | -35.8756
25.6501 | -35.6058 | -35.8809
25.6562 | -35.6062 | -35.8865
25.6623 | -35.6065 | -35.892
25.6684 | -35.6069 | -35.8975
25.6745 | -35.6072 | -35.9029
25.6806 | -35.6075 | -35.9082
25.6867 | -35.6079 | -35.9136
25.6928 | -35.6082 | -35.9193
25.6989 | -35.6085 | -35.9249
25.705 | -35.6089 | -35.9306
25.7111 | -35.6093 | -35.936
25.7173 | -35.6096 | -35.9412
25.7234 | -35.6099 | -35.9468
25.7295 | -35.6103 | -35.9524
25.7356 | -35.6107 | -35.9578
25.7417 | -35.611 | -35.9632
25.7478 | -35.6113 | -35.9686
25.7539 | -35.6116 | -35.9739
25.76 | -35.6119 | -35.9797
25.7661 | -35.6123 | -35.9853
25.7722 | -35.6126 | -35.991
25.7783 | -35.6129 | -35.9965
25.7844 | -35.6132 | -36.0018
25.7905 | -35.6135 | -36.0075
25.7966 | -35.6138 | -36.0128
25.8027 | -35.6142 | -36.0185
25.8088 | -35.6145 | -36.0242
25.8149 | -35.6148 | -36.0295
25.821 | -35.6151 | -36.0347
25.8271 | -35.6154 | -36.0403
25.8332 | -35.6158 | -36.0459
25.8393 | -35.6161 | -36.0514
25.8454 | -35.6165 | -36.0569
25.8515 | -35.6169 | -36.0626
25.8576 | -35.6172 | -36.0681
25.8638 | -35.6175 | -36.0735
25.8699 | -35.6179 | -36.079
25.876 | -35.6182 | -36.0845
25.8821 | -35.6185 | -36.0897
25.8882 | -35.6188 | -36.095
25.8943 | -35.6191 | -36.1005
25.9004 | -35.6194 | -36.1059
25.9065 | -35.6198 | -36.1114
25.9126 | -35.6201 | -36.1169
25.9187 | -35.6204 | -36.1222
25.9248 | -35.6208 | -36.1272
25.9309 | -35.6211 | -36.1327
25.937 | -35.6214 | -36.1382
25.9431 | -35.6218 | -36.1439
25.9492 | -35.6221 | -36.1493
25.9553 | -35.6224 | -36.1549
25.9614 | -35.6227 | -36.16
25.9675 | -35.623 | -36.1654
25.9736 | -35.6233 | -36.1709
25.9797 | -35.6236 | -36.1764
25.9858 | -35.624 | -36.182
25.9919 | -35.6243 | -36.1875
25.998 | -35.6246 | -36.1928
26.0042 | -35.6249 | -36.1983
26.0103 | -35.6252 | -36.2034
26.0164 | -35.6255 | -36.2089
26.0225 | -35.6258 | -36.2141
26.0286 | -35.6261 | -36.2194
26.0347 | -35.6264 | -36.2244
26.0408 | -35.6267 | -36.2297
26.0469 | -35.627 | -36.2348
26.053 | -35.6273 | -36.2404
26.0591 | -35.6277 | -36.2458
26.0652 | -35.6279 | -36.2512
26.0713 | -35.6282 | -36.2569
26.0774 | -35.6285 | -36.2624
26.0835 | -35.6288 | -36.2676
26.0896 | -35.6291 | -36.273
26.0957 | -35.6295 | -36.2781
26.1018 | -35.6298 | -36.2832
26.1079 | -35.6301 | -36.2885
26.114 | -35.6305 | -36.2938
26.1201 | -35.6308 | -36.2991
26.1262 | -35.6312 | -36.3045
26.1323 | -35.6315 | -36.3099
26.1384 | -35.6318 | -36.3151
26.1445 | -35.6321 | -36.3206
26.1507 | -35.6324 | -36.326
26.1568 | -35.6328 | -36.331
26.1629 | -35.6331 | -36.336
26.169 | -35.6334 | -36.3413
26.1751 | -35.6337 | -36.3467
26.1812 | -35.634 | -36.3519
26.1873 | -35.6344 | -36.3574
26.1934 | -35.6348 | -36.3627
26.1995 | -35.6351 | -36.368
26.2056 | -35.6354 | -36.3731
26.2117 | -35.6357 | -36.3786
26.2178 | -35.636 | -36.3838
26.2239 | -35.6363 | -36.3893
26.23 | -35.6366 | -36.3946
26.2361 | -35.6369 | -36.4004
26.2422 | -35.6372 | -36.4056
26.2483 | -35.6375 | -36.4111
26.2544 | -35.6378 | -36.4161
26.2605 | -35.6381 | -36.4215
26.2666 | -35.6384 | -36.4268
26.2727 | -35.6387 | -36.4323
26.2788 | -35.639 | -36.4375
26.2849 | -35.6394 | -36.4428
26.2911 | -35.6397 | -36.4481
26.2972 | -35.64 | -36.4537
26.3033 | -35.6403 | -36.4591
26.3094 | -35.6407 | -36.4643
26.3155 | -35.641 | -36.4693
26.3216 | -35.6413 | -36.4748
26.3277 | -35.6416 | -36.4801
26.3338 | -35.6419 | -36.4853
26.3399 | -35.6423 | -36.4908
26.346 | -35.6426 | -36.4958
26.3521 | -35.6429 | -36.5015
26.3582 | -35.6432 | -36.507
26.3643 | -35.6435 | -36.5125
26.3704 | -35.6438 | -36.518
26.3765 | -35.6441 | -36.5233
26.3826 | -35.6445 | -36.5285
26.3887 | -35.6448 | -36.5337
26.3948 | -35.6451 | -36.539
26.4009 | -35.6454 | -36.5444
26.407 | -35.6457 | -36.5497
26.4131 | -35.646 | -36.5553
26.4192 | -35.6464 | -36.5602
26.4253 | -35.6467 | -36.5656
26.4314 | -35.647 | -36.5709
26.4376 | -35.6474 | -36.5761
26.4437 | -35.6476 | -36.5814
26.4498 | -35.6479 | -36.5868
26.4559 | -35.6483 | -36.5921
26.462 | -35.6486 | -36.5974
26.4681 | -35.6489 | -36.6025
26.4742 | -35.6492 | -36.608
26.4803 | -35.6495 | -36.6128
26.4864 | -35.6498 | -36.6182
26.4925 | -35.6501 | -36.6231
26.4986 | -35.6505 | -36.6281
26.5047 | -35.6508 | -36.6333
26.5108 | -35.6511 | -36.6385
26.5169 | -35.6514 | -36.6439
26.523 | -35.6516 | -36.649
26.5291 | -35.6519 | -36.6542
26.5352 | -35.6522 | -36.6595
26.5413 | -35.6525 | -36.6649
26.5474 | -35.6528 | -36.6701
26.5535 | -35.6531 | -36.6752
26.5596 | -35.6534 | -36.6801
26.5657 | -35.6538 | -36.6854
26.5718 | -35.6541 | -36.6905
26.578 | -35.6544 | -36.6956
26.5841 | -35.6546 | -36.7008
26.5902 | -35.6549 | -36.7063
26.5963 | -35.6552 | -36.7114
26.6024 | -35.6555 | -36.7165
26.6085 | -35.6557 | -36.7216
26.6146 | -35.6561 | -36.7268
26.6207 | -35.6564 | -36.7322
26.6268 | -35.6567 | -36.7375
26.6329 | -35.657 | -36.7423
26.639 | -35.6573 | -36.7474
26.6451 | -35.6576 | -36.7529
26.6512 | -35.6579 | -36.7582
26.6573 | -35.6582 | -36.7633
26.6634 | -35.6585 | -36.768
26.6695 | -35.6588 | -36.7732
26.6756 | -35.6591 | -36.7785
26.6817 | -35.6594 | -36.7835
26.6878 | -35.6596 | -36.7885
26.6939 | -35.66 | -36.7939
26.7 | -35.6603 | -36.7988
26.7061 | -35.6605 | -36.804
26.7122 | -35.6609 | -36.8094
26.7183 | -35.6611 | -36.8149
26.7245 | -35.6613 | -36.8201
26.7306 | -35.6616 | -36.8254
26.7367 | -35.6618 | -36.8302
26.7428 | -35.6621 | -36.8353
26.7489 | -35.6624 | -36.8403
26.755 | -35.6627 | -36.8453
26.7611 | -35.663 | -36.8506
26.7672 | -35.6632 | -36.8557
26.7733 | -35.6635 | -36.8604
26.7794 | -35.6638 | -36.8658
26.7855 | -35.6642 | -36.8711
26.7916 | -35.6645 | -36.8761
26.7977 | -35.6648 | -36.8814
26.8038 | -35.6651 | -36.8864
26.8099 | -35.6654 | -36.8914
26.816 | -35.6657 | -36.8967
26.8221 | -35.666 | -36.9019
26.8282 | -35.6663 | -36.9069
26.8343 | -35.6665 | -36.9121
26.8404 | -35.6668 | -36.9173
26.8465 | -35.6671 | -36.9225
26.8526 | -35.6674 | -36.9278
26.8587 | -35.6678 | -36.9329
26.8649 | -35.6681 | -36.9379
26.871 | -35.6684 | -36.943
26.8771 | -35.6686 | -36.9479
26.8832 | -35.669 | -36.9532
26.8893 | -35.6693 | -36.9582
26.8954 | -35.6695 | -36.9634
26.9015 | -35.6698 | -36.9683
26.9076 | -35.6701 | -36.9737
26.9137 | -35.6704 | -36.9789
26.9198 | -35.6707 | -36.9844
26.9259 | -35.6709 | -36.9893
26.932 | -35.6712 | -36.9944
26.9381 | -35.6715 | -36.9998
26.9442 | -35.6718 | -37.0044
26.9503 | -35.6721 | -37.0096
26.9564 | -35.6724 | -37.015
26.9625 | -35.6727 | -37.0202
26.9686 | -35.673 | -37.0254
26.9747 | -35.6733 | -37.0307
26.9808 | -35.6735 | -37.0356
26.9869 | -35.6738 | -37.0408
26.993 | -35.6741 | -37.0457
26.9991 | -35.6745 | -37.0507
27.0052 | -35.6747 | -37.0557
27.0114 | -35.675 | -37.061
27.0175 | -35.6753 | -37.0661
27.0236 | -35.6756 | -37.071
27.0297 | -35.6758 | -37.076
27.0358 | -35.6761 | -37.081
27.0419 | -35.6765 | -37.0857
27.048 | -35.6767 | -37.0906
27.0541 | -35.677 | -37.0955
27.0602 | -35.6773 | -37.1007
27.0663 | -35.6776 | -37.106
27.0724 | -35.6779 | -37.1112
27.0785 | -35.6782 | -37.1163
27.0846 | -35.6785 | -37.1211
27.0907 | -35.6788 | -37.1262
27.0968 | -35.6791 | -37.1314
27.1029 | -35.6795 | -37.1364
27.109 | -35.6798 | -37.1415
27.1151 | -35.6801 | -37.1468
27.1212 | -35.6804 | -37.1513
27.1273 | -35.6807 | -37.1563
27.1334 | -35.681 | -37.1615
27.1395 | -35.6813 | -37.1666
27.1456 | -35.6816 | -37.1718
27.1518 | -35.6819 | -37.1767
27.1579 | -35.6822 | -37.1816
27.164 | -35.6825 | -37.1866
27.1701 | -35.6828 | -37.1917
27.1762 | -35.6831 | -37.1964
27.1823 | -35.6834 | -37.2016
27.1884 | -35.6837 | -37.2067
27.1945 | -35.684 | -37.2114
27.2006 | -35.6843 | -37.2164
27.2067 | -35.6846 | -37.2216
27.2128 | -35.6849 | -37.2264
27.2189 | -35.6851 | -37.2316
27.225 | -35.6854 | -37.2367
27.2311 | -35.6857 | -37.2416
27.2372 | -35.686 | -37.2467
27.2433 | -35.6863 | -37.2519
27.2494 | -35.6865 | -37.2566
27.2555 | -35.6868 | -37.2618
27.2616 | -35.6871 | -37.2667
27.2677 | -35.6874 | -37.2717
27.2738 | -35.6877 | -37.2767
27.2799 | -35.688 | -37.2818
27.286 | -35.6882 | -37.2867
27.2921 | -35.6885 | -37.2921
27.2983 | -35.6888 | -37.297
27.3044 | -35.6891 | -37.3022
27.3105 | -35.6894 | -37.3073
27.3166 | -35.6897 | -37.3126
27.3227 | -35.69 | -37.3175
27.3288 | -35.6902 | -37.3226
27.3349 | -35.6905 | -37.3277
27.341 | -35.6908 | -37.3331
27.3471 | -35.691 | -37.3382
27.3532 | -35.6913 | -37.3431
27.3593 | -35.6916 | -37.348
27.3654 | -35.6919 | -37.3528
27.3715 | -35.6921 | -37.3579
27.3776 | -35.6924 | -37.363
27.3837 | -35.6927 | -37.3682
27.3898 | -35.693 | -37.373
27.3959 | -35.6932 | -37.378
27.402 | -35.6936 | -37.3828
27.4081 | -35.6939 | -37.3879
27.4142 | -35.6941 | -37.393
27.4203 | -35.6944 | -37.3981
27.4264 | -35.6947 | -37.4031
27.4325 | -35.695 | -37.4083
27.4387 | -35.6954 | -37.413
27.4448 | -35.6956 | -37.4182
27.4509 | -35.6959 | -37.4233
27.457 | -35.6962 | -37.4284
27.4631 | -35.6966 | -37.4338
27.4692 | -35.6968 | -37.4386
27.4753 | -35.697 | -37.4434
27.4814 | -35.6973 | -37.4483
27.4875 | -35.6976 | -37.4531
27.4936 | -35.6979 | -37.4582
27.4997 | -35.6982 | -37.4632
27.5058 | -35.6985 | -37.468
27.5119 | -35.6987 | -37.473
27.518 | -35.699 | -37.4782
27.5241 | -35.6992 | -37.4831
27.5302 | -35.6995 | -37.4882
27.5363 | -35.6997 | -37.4932
27.5424 | -35.7 | -37.4981
27.5485 | -35.7003 | -37.5031
27.5546 | -35.7006 | -37.5081
27.5607 | -35.7009 | -37.513
27.5668 | -35.7012 | -37.5179
27.5729 | -35.7015 | -37.523
27.5791 | -35.7018 | -37.5279
27.5852 | -35.7021 | -37.5327
27.5913 | -35.7023 | -37.5378
27.5974 | -35.7026 | -37.5428
27.6035 | -35.7029 | -37.5479
27.6096 | -35.7031 | -37.5527
27.6157 | -35.7034 | -37.5574
27.6218 | -35.7037 | -37.5626
27.6279 | -35.704 | -37.5676
27.634 | -35.7043 | -37.5725
27.6401 | -35.7046 | -37.5775
27.6462 | -35.7049 | -37.5824
27.6523 | -35.7052 | -37.5871
27.6584 | -35.7055 | -37.5918
27.6645 | -35.7058 | -37.5971
27.6706 | -35.7061 | -37.602
27.6767 | -35.7064 | -37.6072
27.6828 | -35.7066 | -37.612
27.6889 | -35.7069 | -37.617
27.695 | -35.7071 | -37.6219
27.7011 | -35.7074 | -37.6269
27.7072 | -35.7077 | -37.6318
27.7133 | -35.708 | -37.637
27.7194 | -35.7082 | -37.642
27.7256 | -35.7085 | -37.6473
27.7317 | -35.7087 | -37.6522
27.7378 | -35.709 | -37.657
27.7439 | -35.7092 | -37.662
27.75 | -35.7096 | -37.667
27.7561 | -35.7098 | -37.6719
27.7622 | -35.7101 | -37.677
27.7683 | -35.7104 | -37.682
27.7744 | -35.7107 | -37.687
27.7805 | -35.711 | -37.6919
27.7866 | -35.7113 | -37.697
27.7927 | -35.7115 | -37.7015
27.7988 | -35.7118 | -37.7066
27.8049 | -35.7121 | -37.7115
27.811 | -35.7124 | -37.7164
27.8171 | -35.7126 | -37.7214
27.8232 | -35.7129 | -37.7262
27.8293 | -35.7131 | -37.731
27.8354 | -35.7134 | -37.7361
27.8415 | -35.7137 | -37.7409
27.8476 | -35.714 | -37.746
27.8537 | -35.7143 | -37.7507
27.8598 | -35.7145 | -37.7553
27.866 | -35.7148 | -37.7605
27.8721 | -35.7151 | -37.7655
27.8782 | -35.7154 | -37.7705
27.8843 | -35.7156 | -37.7754
27.8904 | -35.7159 | -37.7801
27.8965 | -35.7162 | -37.7846
27.9026 | -35.7164 | -37.7895
27.9087 | -35.7167 | -37.7947
27.9148 | -35.717 | -37.7995
27.9209 | -35.7173 | -37.8045
27.927 | -35.7175 | -37.8093
27.9331 | -35.7178 | -37.8142
27.9392 | -35.718 | -37.8191
27.9453 | -35.7183 | -37.8239
27.9514 | -35.7185 | -37.829
27.9575 | -35.7188 | -37.8342
27.9636 | -35.7191 | -37.8391
27.9697 | -35.7193 | -37.8437
27.9758 | -35.7195 | -37.8488
27.9819 | -35.7198 | -37.8536
27.988 | -35.7201 | -37.8585
27.9941 | -35.7204 | -37.8634
28.0002 | -35.7206 | -37.868
28.0063 | -35.7209 | -37.8731
28.0125 | -35.7212 | -37.8779
28.0186 | -35.7215 | -37.8829
28.0247 | -35.7217 | -37.8878
28.0308 | -35.7219 | -37.8924
28.0369 | -35.7222 | -37.897
28.043 | -35.7225 | -37.9021
28.0491 | -35.7227 | -37.9071
28.0552 | -35.723 | -37.9122
28.0613 | -35.7233 | -37.9171
28.0674 | -35.7236 | -37.9219
28.0735 | -35.7238 | -37.9267

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4586 | -41.6929
21.4908 | -35.5556 | -34.9669
16.4948 | -34.9383 | -27.5988
12.6602 | -34.5821 | -21.3385
9.7171 | -34.3659 | -16.449
7.4582 | -34.2214 | -12.4772
5.7244 | -34.1152 | -9.2654
4.3936 | -34.0535 | -6.665
3.3722 | -34.0334 | -4.6226
2.5883 | -34.0167 | -3.0738
1.9866 | -34.0144 | -1.8332
1.5248 | -34.0232 | -0.83893
1.1703 | -34.0322 | -0.19229
1 | -34.0334 | 0.13002
0.89822 | -34.024 | 0.39901
0.68941 | -34.0188 | 0.98097
0.52919 | -34.0085 | 1.4283
0.40612 | -34.0156 | 2.0328
0.31171 | -34.0167 | 2.4833
0.23928 | -34.0185 | 3.2063
0.18358 | -34.0288 | 3.6253
0.14091 | -34.0456 | 4.5034
0.10819 | -34.068 | 5.6594
0.083017 | -34.086 | 7.1366
0.063705 | -34.1315 | 9.0033
0.048923 | -34.2012 | 11.4808
0.037479 | -34.3115 | 14.6701
0.028801 | -34.4808 | 18.7015
0.022078 | -34.7903 | 24.1435
0.017023 | -35.2797 | 30.5899
0.013018 | -36.087 | 38.2689
0.010014 | -37.2076 | 46.1783

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003628 | 0.47417
0.007256 | -26.9379
0.010884 | -35.9294
0.014512 | -41.2695
0.018141 | -44.5295
0.021769 | -46.3176
0.025397 | -48.2414
0.029025 | -49.9143
0.032653 | -51.51
0.036281 | -52.1424
0.039909 | -52.6987
0.043537 | -54.1842
0.047166 | -54.6838
0.050794 | -55.2873
0.054422 | -56.4049
0.05805 | -57.9305
0.061678 | -59.3277
0.065306 | -59.3697
0.068934 | -59.4822
0.072562 | -60.2126
0.07619 | -62.1094
0.079819 | -62.0221
0.083447 | -62.8127
0.087075 | -63.2545
0.090703 | -63.2127
0.094331 | -63.0867
0.097959 | -63.0655
0.10159 | -63.402
0.10522 | -63.5575
0.10884 | -64.2649
0.11247 | -65.5458
0.1161 | -65.8686
0.11973 | -65.1897
0.12336 | -65.9584
0.12698 | -65.5023
0.13061 | -66.5581
0.13424 | -66.9202
0.13787 | -67.198
0.1415 | -67.6367
0.14513 | -68.3344
0.14875 | -69.1774
0.15238 | -68.72
0.15601 | -69.6163
0.15964 | -69.8539
0.16326 | -70.1857
0.16689 | -69.6352
0.17052 | -69.7461
0.17415 | -70.4192
0.17778 | -70.6354
0.18141 | -70.9489
0.18503 | -71.0804
0.18866 | -71.2396
0.19229 | -71.3624
0.19592 | -70.4243
0.19955 | -71.2992
0.20317 | -71.8616
0.2068 | -72.0562
0.21043 | -71.3587
0.21406 | -71.9985
0.21769 | -72.8035
0.22132 | -72.2451
0.22494 | -72.8334
0.22857 | -73.3432
0.2322 | -73.4492
0.23583 | -72.9077
0.23946 | -72.7003
0.24308 | -73.7979
0.24671 | -73.8443
0.25034 | -74.0839
0.25397 | -73.8606
0.2576 | -74.1094
0.26122 | -74.5479
0.26485 | -75.1528
0.26848 | -75.819
0.27211 | -74.9246
0.27574 | -74.2332
0.27936 | -74.1105
0.28299 | -74.2935
0.28662 | -75.4166
0.29025 | -75.5404
0.29388 | -75.001
0.29751 | -76.1251
0.30113 | -76.3826
0.30476 | -75.4201
0.30839 | -74.6668
0.31202 | -75.4263
0.31565 | -77.0659
0.31927 | -77.5885
0.3229 | -77.5126
0.32653 | -76.7887
0.33016 | -76.7322
0.33379 | -76.0181
0.33742 | -75.9721
0.34104 | -77.2957
0.34467 | -78.2821
0.3483 | -78.1037
0.35193 | -77.4608
0.35556 | -76.8362
0.35918 | -77.3807
0.36281 | -77.6678
0.36644 | -77.2646
0.37007 | -77.0449
0.3737 | -77.5834
0.37732 | -77.7451
0.38095 | -78.3137
0.38458 | -79.0469
0.38821 | -78.8464
0.39184 | -78.437
0.39547 | -77.715
0.39909 | -77.146
0.40272 | -77.1392
0.40635 | -78.6737
0.40998 | -79.46
0.41361 | -79.085
0.41723 | -78.0588
0.42086 | -77.1096
0.42449 | -78.3713
0.42812 | -78.5632
0.43175 | -79.2991
0.43537 | -78.6854
0.439 | -78.4221
0.44263 | -78.6372
0.44626 | -78.7617
0.44989 | -79.8992
0.45352 | -79.0302
0.45714 | -78.439
0.46077 | -78.7142
0.4644 | -79.1618
0.46803 | -79.4195
0.47165 | -79.7149
0.47528 | -79.9876
0.47891 | -80.2993
0.48254 | -79.3258
0.48617 | -78.9322
0.4898 | -79.0305
0.49342 | -79.5349
0.49705 | -79.6431
0.50068 | -79.995
0.50431 | -80.7457
0.50794 | -80.3702
0.51157 | -79.5966
0.51519 | -80.7323
0.51882 | -81.9462
0.52245 | -81.7011
0.52608 | -81.5935
0.5297 | -81.2594
0.53333 | -80.8186
0.53696 | -80.7457
0.54059 | -80.9084
0.54422 | -81.1493
0.54785 | -81.3245
0.55147 | -81.3871
0.5551 | -81.5183
0.55873 | -81.4978
0.56236 | -81.4875
0.56599 | -81.2077
0.56961 | -80.9132
0.57324 | -80.613
0.57687 | -80.3432
0.5805 | -80.715
0.58413 | -81.4516
0.58776 | -81.7887
0.59138 | -81.8278
0.59501 | -81.9561
0.59864 | -81.928
0.60227 | -81.478
0.6059 | -81.3126
0.60952 | -81.5727
0.61315 | -82.0155
0.61678 | -82.6552
0.62041 | -83.011
0.62404 | -82.8095
0.62766 | -82.4888
0.63129 | -82.6343
0.63492 | -82.9896
0.63855 | -83.296
0.64218 | -83.4465
0.6458 | -83.834
0.64943 | -84.1123
0.65306 | -83.6862
0.65669 | -83.2479
0.66032 | -82.9574
0.66395 | -82.8848
0.66757 | -83.0244
0.6712 | -83.1545
0.67483 | -83.3342
0.67846 | -83.147
0.68209 | -82.8945
0.68571 | -83.0187
0.68934 | -83.5656
0.69297 | -84.0173
0.6966 | -84.1654
0.70023 | -84.0663
0.70386 | -83.9391
0.70748 | -84.4008
0.71111 | -85.0825
0.71474 | -85.1919
0.71837 | -84.5407
0.722 | -84.1899
0.72562 | -83.8487
0.72925 | -83.914
0.73288 | -84.3053
0.73651 | -84.4859
0.74014 | -84.7308
0.74376 | -85.3319
0.74739 | -85.575
0.75102 | -85.4523
0.75465 | -84.9799
0.75828 | -84.1126
0.76191 | -83.6107
0.76553 | -83.9288
0.76916 | -84.4228
0.77279 | -84.756
0.77642 | -84.9409
0.78004 | -84.7357
0.78367 | -84.6976
0.7873 | -84.8227
0.79093 | -85.1046
0.79456 | -85.4264
0.79819 | -85.2709
0.80181 | -84.7123
0.80544 | -84.5357
0.80907 | -85.1141
0.8127 | -85.7254
0.81633 | -85.6226
0.81995 | -85.2784
0.82358 | -85.5447
0.82721 | -85.7595
0.83084 | -85.8697
0.83447 | -86.1175
0.8381 | -86.0165
0.84172 | -85.5689
0.84535 | -85.3442
0.84898 | -85.5829
0.85261 | -85.621
0.85624 | -85.3556
0.85986 | -85.3751
0.86349 | -85.6334
0.86712 | -85.472
0.87075 | -85.0142
0.87438 | -84.8454
0.87801 | -85.2797
0.88163 | -86.0436
0.88526 | -86.6859
0.88889 | -86.9671
0.89252 | -86.8865
0.89614 | -86.9296
0.89977 | -86.5973
0.9034 | -86.0519
0.90703 | -85.7395
0.91066 | -85.8028
0.91429 | -86.1544
0.91791 | -86.6796
0.92154 | -86.8679
0.92517 | -86.3387
0.9288 | -85.832
0.93243 | -85.9697
0.93605 | -86.585
0.93968 | -86.9005
0.94331 | -86.7984
0.94694 | -86.4529
0.95057 | -86.2519
0.9542 | -86.4765
0.95782 | -86.8752
0.96145 | -87.0076
0.96508 | -86.9796
0.96871 | -86.7682
0.97234 | -86.7608
0.97596 | -87.0184
0.97959 | -87.2557
0.98322 | -87.3389
0.98685 | -87.3585
0.99048 | -87.3941
0.9941 | -87.7083
0.99773 | -88.0629
1.0014 | -88.023
1.005 | -88.1644
1.0086 | -88.1245
1.0122 | -87.809
1.0159 | -87.4552
1.0195 | -87.3775
1.0231 | -87.5097
1.0268 | -87.8414
1.0304 | -88.1268
1.034 | -88.0033
1.0376 | -87.6759
1.0413 | -87.817
1.0449 | -87.9115
1.0485 | -87.9021
1.0522 | -87.8149
1.0558 | -87.7292
1.0594 | -87.7386
1.063 | -87.831
1.0667 | -87.9618
1.0703 | -88.1306
1.0739 | -88.2272
1.0776 | -88.1846
1.0812 | -88.0838
1.0848 | -87.9959
1.0884 | -87.9751
1.0921 | -88.0598
1.0957 | -88.1528
1.0993 | -88.2239
1.1029 | -88.3895
1.1066 | -88.6477
1.1102 | -88.8327
1.1138 | -88.8288
1.1175 | -88.6605
1.1211 | -88.4218
1.1247 | -88.1998
1.1283 | -88.0922
1.132 | -88.0334
1.1356 | -87.8386
1.1392 | -87.6175
1.1429 | -87.6297
1.1465 | -87.9318
1.1501 | -88.3058
1.1537 | -88.5393
1.1574 | -88.7008
1.161 | -88.7396
1.1646 | -88.4662
1.1683 | -88.0841
1.1719 | -88.001
1.1755 | -88.3269
1.1791 | -88.6763
1.1828 | -88.7203
1.1864 | -88.7492
1.19 | -88.9524
1.1937 | -88.9771
1.1973 | -88.6992
1.2009 | -88.4799
1.2045 | -88.5474
1.2082 | -88.8046
1.2118 | -89.0506
1.2154 | -89.2692
1.219 | -89.4858
1.2227 | -89.4929
1.2263 | -89.2106
1.2299 | -88.9471
1.2336 | -89.0217
1.2372 | -89.3942
1.2408 | -89.7416
1.2444 | -89.8333
1.2481 | -89.7942
1.2517 | -89.7828
1.2553 | -89.6873
1.259 | -89.4334
1.2626 | -89.2085
1.2662 | -89.2731
1.2698 | -89.5951
1.2735 | -89.7563
1.2771 | -89.5517
1.2807 | -89.3106
1.2844 | -89.2848
1.288 | -89.4147
1.2916 | -89.4636
1.2952 | -89.335
1.2989 | -89.2277
1.3025 | -89.3315
1.3061 | -89.5785
1.3098 | -89.7373
1.3134 | -89.7241
1.317 | -89.775
1.3206 | -90.0119
1.3243 | -90.2564
1.3279 | -90.3243
1.3315 | -90.2167
1.3351 | -90.0512
1.3388 | -89.9614
1.3424 | -90.0204
1.346 | -90.1186
1.3497 | -90.0726
1.3533 | -90.0224
1.3569 | -90.1519
1.3605 | -90.2911
1.3642 | -90.1778
1.3678 | -89.8407
1.3714 | -89.4858
1.3751 | -89.3652
1.3787 | -89.6062
1.3823 | -90.0012
1.3859 | -90.2071
1.3896 | -90.2269
1.3932 | -90.3336
1.3968 | -90.6111
1.4005 | -90.7719
1.4041 | -90.5884
1.4077 | -90.2608
1.4113 | -90.0279
1.415 | -89.9016
1.4186 | -89.8042
1.4222 | -89.8108
1.4259 | -90.0002
1.4295 | -90.2358
1.4331 | -90.397
1.4367 | -90.5104
1.4404 | -90.5421
1.444 | -90.5122
1.4476 | -90.4982
1.4512 | -90.5125
1.4549 | -90.548
1.4585 | -90.6311
1.4621 | -90.8223
1.4658 | -91.1666
1.4694 | -91.5922
1.473 | -91.9245
1.4766 | -91.9526
1.4803 | -91.7148
1.4839 | -91.499
1.4875 | -91.4581
1.4912 | -91.5482
1.4948 | -91.6657
1.4984 | -91.7074
1.502 | -91.5755
1.5057 | -91.1461
1.5093 | -90.5104
1.5129 | -90.0725
1.5166 | -90.0102
1.5202 | -90.1106
1.5238 | -90.1884
1.5274 | -90.3241
1.5311 | -90.5935
1.5347 | -90.8393
1.5383 | -90.8606
1.5419 | -90.7118
1.5456 | -90.6218
1.5492 | -90.7102
1.5528 | -90.9133
1.5565 | -91.128
1.5601 | -91.2167
1.5637 | -91.097
1.5673 | -90.9838
1.571 | -90.9439
1.5746 | -90.9744
1.5782 | -91.0534
1.5819 | -91.131
1.5855 | -91.1573
1.5891 | -91.1246
1.5927 | -91.0935
1.5964 | -91.1228
1.6 | -91.2454
1.6036 | -91.442
1.6073 | -91.6609
1.6109 | -91.7909
1.6145 | -91.7507
1.6181 | -91.6006
1.6218 | -91.4888
1.6254 | -91.5043
1.629 | -91.621
1.6327 | -91.731
1.6363 | -91.7431
1.6399 | -91.6656
1.6435 | -91.5607
1.6472 | -91.4419
1.6508 | -91.3046
1.6544 | -91.2167
1.6581 | -91.2448
1.6617 | -91.3921
1.6653 | -91.631
1.6689 | -91.9531
1.6726 | -92.3235
1.6762 | -92.6398
1.6798 | -92.815
1.6834 | -92.8585
1.6871 | -92.8121
1.6907 | -92.6723
1.6943 | -92.4226
1.698 | -92.0629
1.7016 | -91.6519
1.7052 | -91.3438
1.7088 | -91.2816
1.7125 | -91.4233
1.7161 | -91.6311
1.7197 | -91.8114
1.7234 | -91.9814
1.727 | -92.1635
1.7306 | -92.3095
1.7342 | -92.3384
1.7379 | -92.2361
1.7415 | -92.1096
1.7451 | -92.0561
1.7488 | -92.0756
1.7524 | -92.1564
1.756 | -92.263
1.7596 | -92.3821
1.7633 | -92.5269
1.7669 | -92.7408
1.7705 | -93.0298
1.7741 | -93.3381
1.7778 | -93.5308
1.7814 | -93.4452
1.785 | -93.1506
1.7887 | -92.8637
1.7923 | -92.6976
1.7959 | -92.6453
1.7995 | -92.6177
1.8032 | -92.5485
1.8068 | -92.4689
1.8104 | -92.4644
1.8141 | -92.5614
1.8177 | -92.6993
1.8213 | -92.8083
1.8249 | -92.8814
1.8286 | -92.9178
1.8322 | -92.8984
1.8358 | -92.8364
1.8395 | -92.7846
1.8431 | -92.7874
1.8467 | -92.8373
1.8503 | -92.8943
1.854 | -92.8847
1.8576 | -92.7206
1.8612 | -92.4538
1.8649 | -92.2513
1.8685 | -92.2368
1.8721 | -92.3771
1.8757 | -92.4978
1.8794 | -92.4422
1.883 | -92.2821
1.8866 | -92.2149
1.8902 | -92.3243
1.8939 | -92.5812
1.8975 | -92.8968
1.9011 | -93.118
1.9048 | -93.1407
1.9084 | -93.0539
1.912 | -93.0177
1.9156 | -93.0703
1.9193 | -93.1317
1.9229 | -93.1444
1.9265 | -93.1177
1.9302 | -93.1113
1.9338 | -93.1829
1.9374 | -93.3093
1.941 | -93.4283
1.9447 | -93.5147
1.9483 | -93.5993
1.9519 | -93.7243
1.9556 | -93.8587
1.9592 | -93.9258
1.9628 | -93.8814
1.9664 | -93.7297
1.9701 | -93.5371
1.9737 | -93.3765
1.9773 | -93.2815
1.981 | -93.2344
1.9846 | -93.1813
1.9882 | -93.1021
1.9918 | -93.0205
1.9955 | -92.9538
1.9991 | -92.8984
2.0027 | -92.8736
2.0063 | -92.9175
2.01 | -93.0269
2.0136 | -93.1412
2.0172 | -93.2438
2.0209 | -93.3556
2.0245 | -93.4944
2.0281 | -93.6015
2.0317 | -93.6062
2.0354 | -93.4735
2.039 | -93.2602
2.0426 | -93.0859
2.0463 | -93.042
2.0499 | -93.1133
2.0535 | -93.177
2.0571 | -93.1332
2.0608 | -93.0443
2.0644 | -93.0348
2.068 | -93.1588
2.0717 | -93.3605
2.0753 | -93.5706
2.0789 | -93.7831
2.0825 | -94.0224
2.0862 | -94.2923
2.0898 | -94.525
2.0934 | -94.5382
2.0971 | -94.4975
2.1007 | -94.331
2.1043 | -94.0979
2.1079 | -93.8717
2.1116 | -93.7168
2.1152 | -93.6689
2.1188 | -93.7113
2.1224 | -93.7926
2.1261 | -93.8297
2.1297 | -93.7797
2.1333 | -93.6432
2.137 | -93.447
2.1406 | -93.2414
2.1442 | -93.102
2.1478 | -93.0738
2.1515 | -93.1537
2.1551 | -93.2927
2.1587 | -93.4569
2.1624 | -93.6232
2.166 | -93.7771
2.1696 | -93.8881
2.1732 | -93.9539
2.1769 | -93.9646
2.1805 | -93.9223
2.1841 | -93.8277
2.1878 | -93.7105
2.1914 | -93.6105
2.195 | -93.5824
2.1986 | -93.6304
2.2023 | -93.7458
2.2059 | -93.9146
2.2095 | -94.1078
2.2132 | -94.2967
2.2168 | -94.4289
2.2204 | -94.4442
2.224 | -94.3776
2.2277 | -94.2801
2.2313 | -94.2037
2.2349 | -94.1727
2.2385 | -94.2078
2.2422 | -94.2859
2.2458 | -94.36
2.2494 | -94.3971
2.2531 | -94.38
2.2567 | -94.3242
2.2603 | -94.2543
2.2639 | -94.1837
2.2676 | -94.126
2.2712 | -94.0922
2.2748 | -94.0828
2.2785 | -94.0751
2.2821 | -94.0436
2.2857 | -93.9816
2.2893 | -93.9116
2.293 | -93.8878
2.2966 | -93.9218
2.3002 | -93.9645
2.3039 | -93.9847
2.3075 | -93.9825
2.3111 | -93.9878
2.3147 | -94.0516
2.3184 | -94.188
2.322 | -94.4057
2.3256 | -94.6852
2.3293 | -94.9623
2.3329 | -95.1377
2.3365 | -95.1491
2.3401 | -95.0096
2.3438 | -94.7792
2.3474 | -94.5625
2.351 | -94.4083
2.3546 | -94.3186
2.3583 | -94.2947
2.3619 | -94.3357
2.3655 | -94.4628
2.3692 | -94.6626
2.3728 | -94.877
2.3764 | -95.0351
2.38 | -95.0968
2.3837 | -95.0699
2.3873 | -94.9915
2.3909 | -94.9288
2.3946 | -94.9317
2.3982 | -95.0028
2.4018 | -95.0881
2.4054 | -95.1224
2.4091 | -95.0609
2.4127 | -94.9122
2.4163 | -94.7566
2.42 | -94.6499
2.4236 | -94.6133
2.4272 | -94.6317
2.4308 | -94.6978
2.4345 | -94.8207
2.4381 | -94.9736
2.4417 | -95.1099
2.4454 | -95.1967
2.449 | -95.2484
2.4526 | -95.2938
2.4562 | -95.3486
2.4599 | -95.4237
2.4635 | -95.4948
2.4671 | -95.5502
2.4707 | -95.58
2.4744 | -95.5918
2.478 | -95.6061
2.4816 | -95.6051
2.4853 | -95.5577
2.4889 | -95.4648
2.4925 | -95.3516
2.4961 | -95.2354
2.4998 | -95.1151
2.5034 | -94.9872
2.507 | -94.8288
2.5107 | -94.6803
2.5143 | -94.5966
2.5179 | -94.6114
2.5215 | -94.7339
2.5252 | -94.9159
2.5288 | -95.1096
2.5324 | -95.2975
2.5361 | -95.4757
2.5397 | -95.6419
2.5433 | -95.7685
2.5469 | -95.8644
2.5506 | -95.9331
2.5542 | -95.973
2.5578 | -95.9781
2.5615 | -95.9207
2.5651 | -95.8238
2.5687 | -95.683
2.5723 | -95.5008
2.576 | -95.3179
2.5796 | -95.1738
2.5832 | -95.0947
2.5868 | -95.0704
2.5905 | -95.0668
2.5941 | -95.0607
2.5977 | -95.0391
2.6014 | -95.0019
2.605 | -94.9461
2.6086 | -94.8926
2.6122 | -94.8728
2.6159 | -95.0345
2.6195 | -95.1994
2.6231 | -95.4238
2.6268 | -95.6589
2.6304 | -95.831
2.634 | -95.8853
2.6376 | -95.8499
2.6413 | -95.783
2.6449 | -95.7459
2.6485 | -95.7722
2.6522 | -95.8594
2.6558 | -95.9616
2.6594 | -96.0374
2.663 | -96.0486
2.6667 | -95.9942
2.6703 | -95.9008
2.6739 | -95.8042
2.6776 | -95.7427
2.6812 | -95.7425
2.6848 | -95.7864
2.6884 | -95.8507
2.6921 | -95.9104
2.6957 | -95.9499
2.6993 | -95.9717
2.7029 | -95.9915
2.7066 | -96.0033
2.7102 | -96.015
2.7138 | -96.0321
2.7175 | -96.0538
2.7211 | -96.0667
2.7247 | -96.0494
2.7283 | -95.9864
2.732 | -95.8738
2.7356 | -95.7438
2.7392 | -95.6216
2.7429 | -95.544
2.7465 | -95.5289
2.7501 | -95.5875
2.7537 | -95.7143
2.7574 | -95.8915
2.761 | -96.0841
2.7646 | -96.2467
2.7683 | -96.3578
2.7719 | -96.3858
2.7755 | -96.3238
2.7791 | -96.1895
2.7828 | -96.0129
2.7864 | -95.8412
2.79 | -95.717
2.7937 | -95.679
2.7973 | -95.7293
2.8009 | -95.84
2.8045 | -95.9882
2.8082 | -96.1657
2.8118 | -96.3549
2.8154 | -96.5271
2.819 | -96.6697
2.8227 | -96.7756
2.8263 | -96.8395
2.8299 | -96.8541
2.8336 | -96.8087
2.8372 | -96.6953
2.8408 | -96.5018
2.8444 | -96.2509
2.8481 | -95.9916
2.8517 | -95.7828
2.8553 | -95.6935
2.859 | -95.7615
2.8626 | -95.9632
2.8662 | -96.2317
2.8698 | -96.4693
2.8735 | -96.5817
2.8771 | -96.5456
2.8807 | -96.4054
2.8844 | -96.2139
2.888 | -96.0397
2.8916 | -95.9263
2.8952 | -95.9023
2.8989 | -95.9665
2.9025 | -96.0948
2.9061 | -96.2483
2.9098 | -96.3687
2.9134 | -96.4097
2.917 | -96.3686
2.9206 | -96.2732
2.9243 | -96.1515
2.9279 | -96.0333
2.9315 | -95.9405
2.9351 | -95.8856
2.9388 | -95.8894
2.9424 | -95.9665
2.946 | -96.1335
2.9497 | -96.3835
2.9533 | -96.6684
2.9569 | -96.9133
2.9605 | -97.0502
2.9642 | -97.0488
2.9678 | -96.9635
2.9714 | -96.8432
2.9751 | -96.7248
2.9787 | -96.6417
2.9823 | -96.6112
2.9859 | -96.64
2.9896 | -96.7324
2.9932 | -96.8584
2.9968 | -96.9726
3.0005 | -97.0284
3.0041 | -97.0006
3.0077 | -96.9023
3.0113 | -96.7799
3.015 | -96.6554
3.0186 | -96.5308
3.0222 | -96.4048
3.0259 | -96.2819
3.0295 | -96.1932
3.0331 | -96.1654
3.0367 | -96.2037
3.0404 | -96.296
3.044 | -96.4112
3.0476 | -96.5223
3.0512 | -96.5959
3.0549 | -96.5955
3.0585 | -96.5191
3.0621 | -96.4073
3.0658 | -96.3182
3.0694 | -96.2862
3.073 | -96.321
3.0766 | -96.3911
3.0803 | -96.4728
3.0839 | -96.5543
3.0875 | -96.6351
3.0912 | -96.7087
3.0948 | -96.7483
3.0984 | -96.7557
3.102 | -96.7378
3.1057 | -96.7254
3.1093 | -96.7464
3.1129 | -96.8235
3.1166 | -96.9631
3.1202 | -97.1429
3.1238 | -97.3197
3.1274 | -97.4416
3.1311 | -97.5041
3.1347 | -97.5135
3.1383 | -97.494
3.142 | -97.4811
3.1456 | -97.457
3.1492 | -97.4192
3.1528 | -97.3841
3.1565 | -97.3468
3.1601 | -97.3208
3.1637 | -97.3105
3.1673 | -97.3119
3.171 | -97.3151
3.1746 | -97.3388
3.1782 | -97.3617
3.1819 | -97.3609
3.1855 | -97.3603
3.1891 | -97.366
3.1927 | -97.3919
3.1964 | -97.4571
3.2 | -97.5459
3.2036 | -97.643
3.2073 | -97.7382
3.2109 | -97.8174
3.2145 | -97.8686
3.2181 | -97.8954
3.2218 | -97.9083
3.2254 | -97.9218
3.229 | -97.9302
3.2327 | -97.95
3.2363 | -97.966
3.2399 | -97.9465
3.2435 | -97.8849
3.2472 | -97.7693
3.2508 | -97.615
3.2544 | -97.464
3.258 | -97.3392
3.2617 | -97.2495
3.2653 | -97.2023
3.2689 | -97.1777
3.2726 | -97.1515
3.2762 | -97.1048
3.2798 | -97.0366
3.2834 | -96.962
3.2871 | -96.9028
3.2907 | -96.8651
3.2943 | -96.8515
3.298 | -96.8666
3.3016 | -96.9053
3.3052 | -96.9573
3.3088 | -97.0311
3.3125 | -97.1182
3.3161 | -97.2106
3.3197 | -97.2901
3.3234 | -97.3559
3.327 | -97.4141
3.3306 | -97.4705
3.3342 | -97.5402
3.3379 | -97.6269
3.3415 | -97.722
3.3451 | -97.8186
3.3488 | -97.8942
3.3524 | -97.9113
3.356 | -97.8814
3.3596 | -97.8237
3.3633 | -97.7486
3.3669 | -97.6844
3.3705 | -97.6454
3.3742 | -97.6291
3.3778 | -97.6106
3.3814 | -97.588
3.385 | -97.573
3.3887 | -97.5694
3.3923 | -97.5861
3.3959 | -97.6346
3.3995 | -97.6974
3.4032 | -97.7519
3.4068 | -97.7964
3.4104 | -97.8061
3.4141 | -97.7783
3.4177 | -97.7402
3.4213 | -97.7019
3.4249 | -97.6655
3.4286 | -97.6398
3.4322 | -97.6433
3.4358 | -97.6662
3.4395 | -97.6989
3.4431 | -97.7458
3.4467 | -97.8029
3.4503 | -97.8488
3.454 | -97.8826
3.4576 | -97.8978
3.4612 | -97.8857
3.4649 | -97.8384
3.4685 | -97.7722
3.4721 | -97.681
3.4757 | -97.5744
3.4794 | -97.4689
3.483 | -97.3806
3.4866 | -97.3295
3.4902 | -97.3225
3.4939 | -97.3473
3.4975 | -97.3803
3.5011 | -97.4142
3.5048 | -97.4646
3.5084 | -97.5247
3.512 | -97.5917
3.5156 | -97.6664
3.5193 | -97.7474
3.5229 | -97.8211
3.5265 | -97.8863
3.5302 | -97.9445
3.5338 | -97.9984
3.5374 | -98.0498
3.541 | -98.0994
3.5447 | -98.1635
3.5483 | -98.2272
3.5519 | -98.2791
3.5556 | -98.3015
3.5592 | -98.2845
3.5628 | -98.2327
3.5664 | -98.1678
3.5701 | -98.0857
3.5737 | -98.0129
3.5773 | -97.9679
3.581 | -97.9761
3.5846 | -98.0422
3.5882 | -98.1541
3.5918 | -98.3078
3.5955 | -98.474
3.5991 | -98.6214
3.6027 | -98.739
3.6063 | -98.7934
3.61 | -98.7741
3.6136 | -98.6796
3.6172 | -98.5481
3.6209 | -98.3941
3.6245 | -98.2793
3.6281 | -98.2221
3.6317 | -98.2362
3.6354 | -98.3144
3.639 | -98.4342
3.6426 | -98.5761
3.6463 | -98.6983
3.6499 | -98.7863
3.6535 | -98.8335
3.6571 | -98.8329
3.6608 | -98.8106
3.6644 | -98.8066
3.668 | -98.7902
3.6717 | -98.7805
3.6753 | -98.766
3.6789 | -98.749
3.6825 | -98.7225
3.6862 | -98.6963
3.6898 | -98.668
3.6934 | -98.6471
3.6971 | -98.6401
3.7007 | -98.6423
3.7043 | -98.6632
3.7079 | -98.7017
3.7116 | -98.7525
3.7152 | -98.8085
3.7188 | -98.8606
3.7224 | -98.8945
3.7261 | -98.9258
3.7297 | -98.9451
3.7333 | -98.9303
3.737 | -98.9019
3.7406 | -98.8557
3.7442 | -98.8009
3.7478 | -98.7541
3.7515 | -98.7298
3.7551 | -98.7273
3.7587 | -98.7485
3.7624 | -98.785
3.766 | -98.8203
3.7696 | -98.8466
3.7732 | -98.868
3.7769 | -98.8893
3.7805 | -98.9129
3.7841 | -98.9496
3.7878 | -98.9819
3.7914 | -98.9996
3.795 | -98.9968
3.7986 | -98.9501
3.8023 | -98.8689
3.8059 | -98.7682
3.8095 | -98.6531
3.8132 | -98.5416
3.8168 | -98.4501
3.8204 | -98.3799
3.824 | -98.3373
3.8277 | -98.3346
3.8313 | -98.3622
3.8349 | -98.4037
3.8385 | -98.4439
3.8422 | -98.4789
3.8458 | -98.5096
3.8494 | -98.5292
3.8531 | -98.5464
3.8567 | -98.5784
3.8603 | -98.6241
3.8639 | -98.6825
3.8676 | -98.7531
3.8712 | -98.8385
3.8748 | -98.9199
3.8785 | -98.9908
3.8821 | -99.0475
3.8857 | -99.0932
3.8893 | -99.1238
3.893 | -99.1456
3.8966 | -99.1436
3.9002 | -99.119
3.9039 | -99.0656
3.9075 | -98.9787
3.9111 | -98.8884
3.9147 | -98.8067
3.9184 | -98.7332
3.922 | -98.6818
3.9256 | -98.6421
3.9293 | -98.6035
3.9329 | -98.5788
3.9365 | -98.5836
3.9401 | -98.6123
3.9438 | -98.6701
3.9474 | -98.7455
3.951 | -98.8277
3.9546 | -98.9251
3.9583 | -99.0411
3.9619 | -99.1606
3.9655 | -99.2758
3.9692 | -99.3984
3.9728 | -99.5124
3.9764 | -99.6129
3.98 | -99.7027
3.9837 | -99.7909
3.9873 | -99.8672
3.9909 | -99.9367
3.9946 | -99.9604
3.9982 | -99.9053
4.0018 | -99.7884
4.0054 | -99.6121
4.0091 | -99.4
4.0127 | -99.1755
4.0163 | -98.9595
4.02 | -98.7729
4.0236 | -98.6218
4.0272 | -98.5221
4.0308 | -98.4793
4.0345 | -98.4891
4.0381 | -98.5462
4.0417 | -98.6513
4.0454 | -98.7786
4.049 | -98.9099
4.0526 | -99.0176
4.0562 | -99.1157
4.0599 | -99.2038
4.0635 | -99.2626
4.0671 | -99.2961
4.0707 | -99.3135
4.0744 | -99.3282
4.078 | -99.3392
4.0816 | -99.3635
4.0853 | -99.3835
4.0889 | -99.3861
4.0925 | -99.3558
4.0961 | -99.2974
4.0998 | -99.2255
4.1034 | -99.153
4.107 | -99.1122
4.1107 | -99.1266
4.1143 | -99.1942
4.1179 | -99.2985
4.1215 | -99.4401
4.1252 | -99.6049
4.1288 | -99.7714
4.1324 | -99.9062
4.1361 | -99.9968
4.1397 | -100.0411
4.1433 | -100.0288
4.1469 | -99.9709
4.1506 | -99.8758
4.1542 | -99.7729
4.1578 | -99.6798
4.1615 | -99.6015
4.1651 | -99.5539
4.1687 | -99.5303
4.1723 | -99.5249
4.176 | -99.5344
4.1796 | -99.5493
4.1832 | -99.5576
4.1868 | -99.5372
4.1905 | -99.5232
4.1941 | -99.5006
4.1977 | -99.4717
4.2014 | -99.4327
4.205 | -99.3895
4.2086 | -99.3475
4.2122 | -99.3177
4.2159 | -99.2949
4.2195 | -99.2713
4.2231 | -99.2454
4.2268 | -99.2104
4.2304 | -99.1644
4.234 | -99.1248
4.2376 | -99.0862
4.2413 | -99.0607
4.2449 | -99.0609
4.2485 | -99.0872
4.2522 | -99.1356
4.2558 | -99.2063
4.2594 | -99.2974
4.263 | -99.4048
4.2667 | -99.5236
4.2703 | -99.6483
4.2739 | -99.771
4.2776 | -99.8855
4.2812 | -99.9663
4.2848 | -100.0141
4.2884 | -100.0387
4.2921 | -100.0316
4.2957 | -100.0018
4.2993 | -99.9634
4.3029 | -99.9153
4.3066 | -99.8648
4.3102 | -99.8095
4.3138 | -99.7549
4.3175 | -99.7103
4.3211 | -99.6716
4.3247 | -99.6432
4.3283 | -99.6326
4.332 | -99.6259
4.3356 | -99.6205
4.3392 | -99.6078
4.3429 | -99.5914
4.3465 | -99.5681
4.3501 | -99.5492
4.3537 | -99.5328
4.3574 | -99.523
4.361 | -99.5235
4.3646 | -99.5361
4.3683 | -99.5556
4.3719 | -99.595
4.3755 | -99.6519
4.3791 | -99.725
4.3828 | -99.8094
4.3864 | -99.8951
4.39 | -99.9877
4.3937 | -100.0623
4.3973 | -100.1049
4.4009 | -100.1292
4.4045 | -100.1302
4.4082 | -100.11
4.4118 | -100.0711
4.4154 | -100.0365
4.419 | -100.0105
4.4227 | -99.994
4.4263 | -99.9896
4.4299 | -99.9973
4.4336 | -100.0151
4.4372 | -100.0398
4.4408 | -100.0681
4.4444 | -100.1095
4.4481 | -100.1546
4.4517 | -100.2003
4.4553 | -100.2529
4.459 | -100.3005
4.4626 | -100.3335
4.4662 | -100.3432
4.4698 | -100.3225
4.4735 | -100.2779
4.4771 | -100.2217
4.4807 | -100.1584
4.4844 | -100.0924
4.488 | -100.0245
4.4916 | -99.9591
4.4952 | -99.902
4.4989 | -99.859
4.5025 | -99.8224
4.5061 | -99.7981
4.5098 | -99.7932
4.5134 | -99.8001
4.517 | -99.8273
4.5206 | -99.8595
4.5243 | -99.8893
4.5279 | -99.9308
4.5315 | -99.9819
4.5351 | -100.0428
4.5388 | -100.106
4.5424 | -100.1858
4.546 | -100.27
4.5497 | -100.357
4.5533 | -100.4404
4.5569 | -100.5143
4.5605 | -100.5648
4.5642 | -100.5873
4.5678 | -100.5759
4.5714 | -100.5397
4.5751 | -100.4791
4.5787 | -100.4006
4.5823 | -100.3192
4.5859 | -100.2342
4.5896 | -100.1585
4.5932 | -100.0946
4.5968 | -100.0415
4.6005 | -100.0055
4.6041 | -99.9808
4.6077 | -99.9661
4.6113 | -99.9577
4.615 | -99.9634
4.6186 | -99.9712
4.6222 | -99.974
4.6258 | -99.9792
4.6295 | -99.9788
4.6331 | -99.9761
4.6367 | -99.9783
4.6404 | -99.9978
4.644 | -100.0344
4.6476 | -100.0879
4.6512 | -100.1575
4.6549 | -100.2443
4.6585 | -100.3451
4.6621 | -100.444
4.6658 | -100.5306
4.6694 | -100.6067
4.673 | -100.6764
4.6766 | -100.7298
4.6803 | -100.7598
4.6839 | -100.775
4.6875 | -100.7793
4.6912 | -100.779
4.6948 | -100.7711
4.6984 | -100.7683
4.702 | -100.771
4.7057 | -100.7826
4.7093 | -100.8023
4.7129 | -100.8088
4.7166 | -100.8117
4.7202 | -100.805
4.7238 | -100.78
4.7274 | -100.7279
4.7311 | -100.6561
4.7347 | -100.5776
4.7383 | -100.4972
4.742 | -100.418
4.7456 | -100.3293
4.7492 | -100.2478
4.7528 | -100.1725
4.7565 | -100.1122
4.7601 | -100.0741
4.7637 | -100.0567
4.7673 | -100.0687
4.771 | -100.1095
4.7746 | -100.1669
4.7782 | -100.242
4.7819 | -100.3283
4.7855 | -100.4193
4.7891 | -100.5095
4.7927 | -100.5973
4.7964 | -100.6698
4.8 | -100.7318
4.8036 | -100.7738
4.8073 | -100.7931
4.8109 | -100.7921
4.8145 | -100.7707
4.8181 | -100.7388
4.8218 | -100.7051
4.8254 | -100.6649
4.829 | -100.63
4.8327 | -100.6065
4.8363 | -100.5909
4.8399 | -100.5894
4.8435 | -100.5945
4.8472 | -100.6094
4.8508 | -100.6333
4.8544 | -100.656
4.8581 | -100.6695
4.8617 | -100.6677
4.8653 | -100.6599
4.8689 | -100.6443
4.8726 | -100.6248
4.8762 | -100.5988
4.8798 | -100.5775
4.8834 | -100.5627
4.8871 | -100.5544
4.8907 | -100.5551
4.8943 | -100.5668
4.898 | -100.5962
4.9016 | -100.643
4.9052 | -100.6979
4.9088 | -100.7507
4.9125 | -100.8041
4.9161 | -100.8504
4.9197 | -100.8834
4.9234 | -100.8985
4.927 | -100.9019
4.9306 | -100.8825
4.9342 | -100.8542
4.9379 | -100.8197
4.9415 | -100.7686
4.9451 | -100.7171
4.9488 | -100.6826
4.9524 | -100.6636
4.956 | -100.6623
4.9596 | -100.678
4.9633 | -100.7178
4.9669 | -100.7761
4.9705 | -100.8411
4.9741 | -100.9023
4.9778 | -100.9574
4.9814 | -100.9981
4.985 | -101.0278
4.9887 | -101.0435
4.9923 | -101.0406
4.9959 | -101.0294
4.9995 | -101.0101
5.0032 | -100.9906
5.0068 | -100.977
5.0104 | -100.9699
5.0141 | -100.9718
5.0177 | -100.9728
5.0213 | -100.9654
5.0249 | -100.9498
5.0286 | -100.9255
5.0322 | -100.8925
5.0358 | -100.8591
5.0395 | -100.8208
5.0431 | -100.7941
5.0467 | -100.7807
5.0503 | -100.7813
5.054 | -100.7892
5.0576 | -100.8005
5.0612 | -100.8168
5.0649 | -100.8285
5.0685 | -100.8426
5.0721 | -100.8521
5.0757 | -100.8537
5.0794 | -100.8469
5.083 | -100.8228
5.0866 | -100.7791
5.0902 | -100.7216
5.0939 | -100.6611
5.0975 | -100.5965
5.1011 | -100.5235
5.1048 | -100.4692
5.1084 | -100.4304
5.112 | -100.41
5.1156 | -100.4092
5.1193 | -100.4239
5.1229 | -100.4546
5.1265 | -100.4949
5.1302 | -100.5443
5.1338 | -100.6106
5.1374 | -100.6798
5.141 | -100.7601
5.1447 | -100.844
5.1483 | -100.9296
5.1519 | -101.0196
5.1556 | -101.0965
5.1592 | -101.173
5.1628 | -101.2359
5.1664 | -101.2769
5.1701 | -101.3061
5.1737 | -101.3257
5.1773 | -101.3486
5.181 | -101.3453
5.1846 | -101.3331
5.1882 | -101.3121
5.1918 | -101.2831
5.1955 | -101.2432
5.1991 | -101.1946
5.2027 | -101.1532
5.2063 | -101.1186
5.21 | -101.0866
5.2136 | -101.0686
5.2172 | -101.0582
5.2209 | -101.0581
5.2245 | -101.0668
5.2281 | -101.0735
5.2317 | -101.0812
5.2354 | -101.0942
5.239 | -101.1053
5.2426 | -101.121
5.2463 | -101.1478
5.2499 | -101.1785
5.2535 | -101.2088
5.2571 | -101.2354
5.2608 | -101.2679
5.2644 | -101.2965
5.268 | -101.3134
5.2717 | -101.3274
5.2753 | -101.3428
5.2789 | -101.3571
5.2825 | -101.3697
5.2862 | -101.3951
5.2898 | -101.437
5.2934 | -101.4892
5.2971 | -101.5438
5.3007 | -101.5899
5.3043 | -101.6452
5.3079 | -101.6955
5.3116 | -101.7377
5.3152 | -101.7611
5.3188 | -101.7621
5.3224 | -101.7508
5.3261 | -101.72
5.3297 | -101.6814
5.3333 | -101.6383
5.337 | -101.5959
5.3406 | -101.5602
5.3442 | -101.53
5.3478 | -101.4961
5.3515 | -101.4784
5.3551 | -101.4662
5.3587 | -101.4511
5.3624 | -101.4356
5.366 | -101.4153
5.3696 | -101.392
5.3732 | -101.3715
5.3769 | -101.3603
5.3805 | -101.3652
5.3841 | -101.3791
5.3878 | -101.3939
5.3914 | -101.4198
5.395 | -101.4485
5.3986 | -101.4864
5.4023 | -101.5236
5.4059 | -101.5516
5.4095 | -101.5763
5.4132 | -101.5863
5.4168 | -101.58
5.4204 | -101.5656
5.424 | -101.5447
5.4277 | -101.5153
5.4313 | -101.4946
5.4349 | -101.4685
5.4385 | -101.4435
5.4422 | -101.4319
5.4458 | -101.4356
5.4494 | -101.4492
5.4531 | -101.4586
5.4567 | -101.4589
5.4603 | -101.4616
5.4639 | -101.4601
5.4676 | -101.456
5.4712 | -101.4465
5.4748 | -101.4386
5.4785 | -101.4304
5.4821 | -101.4183
5.4857 | -101.4058
5.4893 | -101.4053
5.493 | -101.4154
5.4966 | -101.4304
5.5002 | -101.4486
5.5039 | -101.4653
5.5075 | -101.4747
5.5111 | -101.4797
5.5147 | -101.4689
5.5184 | -101.448
5.522 | -101.4137
5.5256 | -101.3728
5.5293 | -101.3217
5.5329 | -101.2708
5.5365 | -101.2205
5.5401 | -101.1794
5.5438 | -101.1384
5.5474 | -101.1029
5.551 | -101.0741
5.5546 | -101.0491
5.5583 | -101.0277
5.5619 | -101.0202
5.5655 | -101.0217
5.5692 | -101.0302
5.5728 | -101.0395
5.5764 | -101.0457
5.58 | -101.0597
5.5837 | -101.0765
5.5873 | -101.0954
5.5909 | -101.1196
5.5946 | -101.1452
5.5982 | -101.1687
5.6018 | -101.1889
5.6054 | -101.2125
5.6091 | -101.2247
5.6127 | -101.235
5.6163 | -101.248
5.62 | -101.2644
5.6236 | -101.2642
5.6272 | -101.2526
5.6308 | -101.2383
5.6345 | -101.2174
5.6381 | -101.1962
5.6417 | -101.1759
5.6454 | -101.1503
5.649 | -101.1302
5.6526 | -101.1195
5.6562 | -101.1161
5.6599 | -101.1174
5.6635 | -101.1124
5.6671 | -101.1074
5.6707 | -101.0944
5.6744 | -101.0666
5.678 | -101.0347
5.6816 | -100.9975
5.6853 | -100.9608
5.6889 | -100.9207
5.6925 | -100.8827
5.6961 | -100.85
5.6998 | -100.8197
5.7034 | -100.8058
5.707 | -100.808
5.7107 | -100.8136
5.7143 | -100.8136
5.7179 | -100.8192
5.7215 | -100.8249
5.7252 | -100.8312
5.7288 | -100.839
5.7324 | -100.8403
5.7361 | -100.8381
5.7397 | -100.8321
5.7433 | -100.8332
5.7469 | -100.8436
5.7506 | -100.8542
5.7542 | -100.8737
5.7578 | -100.8948
5.7615 | -100.921
5.7651 | -100.9452
5.7687 | -100.9711
5.7723 | -100.9992
5.776 | -101.0312
5.7796 | -101.0623
5.7832 | -101.0926
5.7868 | -101.1291
5.7905 | -101.1747
5.7941 | -101.2307
5.7977 | -101.2935
5.8014 | -101.3626
5.805 | -101.4204
5.8086 | -101.468
5.8122 | -101.5106
5.8159 | -101.5327
5.8195 | -101.5348
5.8231 | -101.5236
5.8268 | -101.4997
5.8304 | -101.4632
5.834 | -101.4152
5.8376 | -101.3592
5.8413 | -101.3107
5.8449 | -101.2648
5.8485 | -101.2215
5.8522 | -101.1924
5.8558 | -101.1721
5.8594 | -101.1633
5.863 | -101.1595
5.8667 | -101.1624
5.8703 | -101.1768
5.8739 | -101.2016
5.8776 | -101.2245
5.8812 | -101.2373
5.8848 | -101.2512
5.8884 | -101.265
5.8921 | -101.2712
5.8957 | -101.2745
5.8993 | -101.2777
5.9029 | -101.2719
5.9066 | -101.2738
5.9102 | -101.2772
5.9138 | -101.2745
5.9175 | -101.269
5.9211 | -101.2636
5.9247 | -101.2621
5.9283 | -101.2631
5.932 | -101.26
5.9356 | -101.2611
5.9392 | -101.2664
5.9429 | -101.2764
5.9465 | -101.3007
5.9501 | -101.3306
5.9537 | -101.3677
5.9574 | -101.4179
5.961 | -101.4736
5.9646 | -101.5292
5.9683 | -101.5896
5.9719 | -101.6508
5.9755 | -101.7188
5.9791 | -101.7956
5.9828 | -101.8678
5.9864 | -101.938
5.99 | -102.003
5.9937 | -102.0531
5.9973 | -102.0883
6.0009 | -102.1094
6.0045 | -102.1145
6.0082 | -102.1029
6.0118 | -102.0763
6.0154 | -102.0368
6.019 | -101.9983
6.0227 | -101.9565
6.0263 | -101.9161
6.0299 | -101.885
6.0336 | -101.866
6.0372 | -101.8511
6.0408 | -101.844
6.0444 | -101.8425
6.0481 | -101.8402
6.0517 | -101.8358
6.0553 | -101.8358
6.059 | -101.8362
6.0626 | -101.8368
6.0662 | -101.8311
6.0698 | -101.8179
6.0735 | -101.8117
6.0771 | -101.8041
6.0807 | -101.7922
6.0844 | -101.7885
6.088 | -101.7906
6.0916 | -101.8035
6.0952 | -101.8266
6.0989 | -101.8636
6.1025 | -101.9106
6.1061 | -101.9679
6.1098 | -102.027
6.1134 | -102.0833
6.117 | -102.134
6.1206 | -102.1768
6.1243 | -102.2014
6.1279 | -102.2119
6.1315 | -102.2069
6.1351 | -102.1874
6.1388 | -102.1598
6.1424 | -102.1314
6.146 | -102.1018
6.1497 | -102.0822
6.1533 | -102.072
6.1569 | -102.0698
6.1605 | -102.084
6.1642 | -102.1145
6.1678 | -102.1596
6.1714 | -102.2138
6.1751 | -102.2736
6.1787 | -102.3309
6.1823 | -102.3838
6.1859 | -102.432
6.1896 | -102.4785
6.1932 | -102.5189
6.1968 | -102.5397
6.2005 | -102.5501
6.2041 | -102.5521
6.2077 | -102.5361
6.2113 | -102.495
6.215 | -102.4453
6.2186 | -102.3779
6.2222 | -102.2995
6.2259 | -102.2238
6.2295 | -102.1445
6.2331 | -102.068
6.2367 | -101.9985
6.2404 | -101.9384
6.244 | -101.8927
6.2476 | -101.862
6.2512 | -101.8463
6.2549 | -101.8421
6.2585 | -101.853
6.2621 | -101.8828
6.2658 | -101.9257
6.2694 | -101.9832
6.273 | -102.0512
6.2766 | -102.1216
6.2803 | -102.177
6.2839 | -102.2308
6.2875 | -102.2821
6.2912 | -102.33
6.2948 | -102.3635
6.2984 | -102.3852
6.302 | -102.3946
6.3057 | -102.3953
6.3093 | -102.3865
6.3129 | -102.3686
6.3166 | -102.344
6.3202 | -102.3091
6.3238 | -102.274
6.3274 | -102.2461
6.3311 | -102.219
6.3347 | -102.1952
6.3383 | -102.1767
6.3419 | -102.1603
6.3456 | -102.1483
6.3492 | -102.1448
6.3528 | -102.1437
6.3565 | -102.1427
6.3601 | -102.1381
6.3637 | -102.1349
6.3673 | -102.1377
6.371 | -102.1458
6.3746 | -102.1592
6.3782 | -102.1755
6.3819 | -102.1971
6.3855 | -102.2177
6.3891 | -102.2415
6.3927 | -102.2673
6.3964 | -102.2865
6.4 | -102.3058
6.4036 | -102.3291
6.4073 | -102.353
6.4109 | -102.371
6.4145 | -102.384
6.4181 | -102.4017
6.4218 | -102.4194
6.4254 | -102.4328
6.429 | -102.4512
6.4327 | -102.4709
6.4363 | -102.4909
6.4399 | -102.5084
6.4435 | -102.5248
6.4472 | -102.5355
6.4508 | -102.5481
6.4544 | -102.5647
6.4581 | -102.5862
6.4617 | -102.6002
6.4653 | -102.6108
6.4689 | -102.6205
6.4726 | -102.6272
6.4762 | -102.6278
6.4798 | -102.6154
6.4834 | -102.602
6.4871 | -102.5866
6.4907 | -102.5604
6.4943 | -102.5314
6.498 | -102.4988
6.5016 | -102.4616
6.5052 | -102.4223
6.5088 | -102.3898
6.5125 | -102.3589
6.5161 | -102.3243
6.5197 | -102.2879
6.5234 | -102.2444
6.527 | -102.2081
6.5306 | -102.1787
6.5342 | -102.158
6.5379 | -102.152
6.5415 | -102.1511
6.5451 | -102.1554
6.5488 | -102.1701
6.5524 | -102.1958
6.556 | -102.2205
6.5596 | -102.2547
6.5633 | -102.2936
6.5669 | -102.3299
6.5705 | -102.3645
6.5742 | -102.3919
6.5778 | -102.4222
6.5814 | -102.4496
6.585 | -102.467
6.5887 | -102.4847
6.5923 | -102.5027
6.5959 | -102.5143
6.5995 | -102.5226
6.6032 | -102.5295
6.6068 | -102.5387
6.6104 | -102.5503
6.6141 | -102.5651
6.6177 | -102.5843
6.6213 | -102.6139
6.6249 | -102.6359
6.6286 | -102.6557
6.6322 | -102.6743
6.6358 | -102.6912
6.6395 | -102.7082
6.6431 | -102.7319
6.6467 | -102.7465
6.6503 | -102.7531
6.654 | -102.7651
6.6576 | -102.7716
6.6612 | -102.7717
6.6649 | -102.773
6.6685 | -102.7673
6.6721 | -102.7543
6.6757 | -102.7355
6.6794 | -102.7115
6.683 | -102.6909
6.6866 | -102.6646
6.6902 | -102.6304
6.6939 | -102.5906
6.6975 | -102.5507
6.7011 | -102.5116
6.7048 | -102.4764
6.7084 | -102.4441
6.712 | -102.4172
6.7156 | -102.4018
6.7193 | -102.3943
6.7229 | -102.3993
6.7265 | -102.4179
6.7302 | -102.4484
6.7338 | -102.4928
6.7374 | -102.5502
6.741 | -102.617
6.7447 | -102.6938
6.7483 | -102.777
6.7519 | -102.8619
6.7556 | -102.9444
6.7592 | -103.0298
6.7628 | -103.109
6.7664 | -103.1797
6.7701 | -103.2388
6.7737 | -103.2823
6.7773 | -103.3196
6.781 | -103.345
6.7846 | -103.3653
6.7882 | -103.3813
6.7918 | -103.3948
6.7955 | -103.4037
6.7991 | -103.403
6.8027 | -103.3906
6.8063 | -103.3786
6.81 | -103.363
6.8136 | -103.3458
6.8172 | -103.3273
6.8209 | -103.3093
6.8245 | -103.2867
6.8281 | -103.2635
6.8317 | -103.2389
6.8354 | -103.2111
6.839 | -103.1859
6.8426 | -103.1605
6.8463 | -103.1402
6.8499 | -103.1188
6.8535 | -103.0962
6.8571 | -103.0766
6.8608 | -103.0518
6.8644 | -103.0318
6.868 | -103.0155
6.8717 | -103.0023
6.8753 | -102.9824
6.8789 | -102.9611
6.8825 | -102.9416
6.8862 | -102.9274
6.8898 | -102.9179
6.8934 | -102.9125
6.8971 | -102.9116
6.9007 | -102.9209
6.9043 | -102.9338
6.9079 | -102.9521
6.9116 | -102.9766
6.9152 | -103.0036
6.9188 | -103.0351
6.9224 | -103.0699
6.9261 | -103.1004
6.9297 | -103.1352
6.9333 | -103.1624
6.937 | -103.1878
6.9406 | -103.2093
6.9442 | -103.2214
6.9478 | -103.221
6.9515 | -103.2178
6.9551 | -103.2154
6.9587 | -103.2087
6.9624 | -103.2005
6.966 | -103.1948
6.9696 | -103.1854
6.9732 | -103.1748
6.9769 | -103.1705
6.9805 | -103.1761
6.9841 | -103.1816
6.9878 | -103.1899
6.9914 | -103.1963
6.995 | -103.2009
6.9986 | -103.2048
7.0023 | -103.1969
7.0059 | -103.1899
7.0095 | -103.1789
7.0132 | -103.1644
7.0168 | -103.1534
7.0204 | -103.1393
7.024 | -103.1187
7.0277 | -103.0991
7.0313 | -103.0814
7.0349 | -103.0619
7.0385 | -103.0465
7.0422 | -103.0331
7.0458 | -103.0235
7.0494 | -103.0164
7.0531 | -103.015
7.0567 | -103.0176
7.0603 | -103.0283
7.0639 | -103.0497
7.0676 | -103.0729
7.0712 | -103.1053
7.0748 | -103.1464
7.0785 | -103.1867
7.0821 | -103.2194
7.0857 | -103.2492
7.0893 | -103.2768
7.093 | -103.2905
7.0966 | -103.2974
7.1002 | -103.2983
7.1039 | -103.2905
7.1075 | -103.2853
7.1111 | -103.274
7.1147 | -103.2663
7.1184 | -103.2584
7.122 | -103.2472
7.1256 | -103.2437
7.1293 | -103.2483
7.1329 | -103.2487
7.1365 | -103.2532
7.1401 | -103.2621
7.1438 | -103.272
7.1474 | -103.2773
7.151 | -103.278
7.1546 | -103.2743
7.1583 | -103.2643
7.1619 | -103.2475
7.1655 | -103.2274
7.1692 | -103.2021
7.1728 | -103.179
7.1764 | -103.1426
7.18 | -103.0994
7.1837 | -103.065
7.1873 | -103.0389
7.1909 | -103.0115
7.1946 | -102.9949
7.1982 | -102.9936
7.2018 | -103.0002
7.2054 | -103.0176
7.2091 | -103.0472
7.2127 | -103.0889
7.2163 | -103.1387
7.22 | -103.1916
7.2236 | -103.2452
7.2272 | -103.3023
7.2308 | -103.3595
7.2345 | -103.4087
7.2381 | -103.4585
7.2417 | -103.5042
7.2454 | -103.5375
7.249 | -103.5592
7.2526 | -103.5788
7.2562 | -103.5947
7.2599 | -103.6082
7.2635 | -103.6152
7.2671 | -103.6131
7.2707 | -103.6101
7.2744 | -103.6072
7.278 | -103.6052
7.2816 | -103.6164
7.2853 | -103.6325
7.2889 | -103.6448
7.2925 | -103.6526
7.2961 | -103.6717
7.2998 | -103.6939
7.3034 | -103.7227
7.307 | -103.7526
7.3107 | -103.786
7.3143 | -103.8166
7.3179 | -103.8415
7.3215 | -103.8595
7.3252 | -103.8417
7.3288 | -103.8393
7.3324 | -103.8338
7.3361 | -103.8172
7.3397 | -103.7921
7.3433 | -103.7601
7.3469 | -103.7227
7.3506 | -103.6851
7.3542 | -103.6502
7.3578 | -103.62
7.3615 | -103.5903
7.3651 | -103.5608
7.3687 | -103.5298
7.3723 | -103.5057
7.376 | -103.4837
7.3796 | -103.4576
7.3832 | -103.4339
7.3868 | -103.4057
7.3905 | -103.3751
7.3941 | -103.3483
7.3977 | -103.3172
7.4014 | -103.284
7.405 | -103.25
7.4086 | -103.2177
7.4122 | -103.1878
7.4159 | -103.1625
7.4195 | -103.1439
7.4231 | -103.1339
7.4268 | -103.1276
7.4304 | -103.1348
7.434 | -103.1544
7.4376 | -103.1835
7.4413 | -103.222
7.4449 | -103.2632
7.4485 | -103.3112
7.4522 | -103.3602
7.4558 | -103.4164
7.4594 | -103.4746
7.463 | -103.5282
7.4667 | -103.5825
7.4703 | -103.6365
7.4739 | -103.6865
7.4776 | -103.7356
7.4812 | -103.7862
7.4848 | -103.8295
7.4884 | -103.8677
7.4921 | -103.9106
7.4957 | -103.9489
7.4993 | -103.9877
7.5029 | -104.0207
7.5066 | -104.0386
7.5102 | -104.0583
7.5138 | -104.0682
7.5175 | -104.0663
7.5211 | -104.0565
7.5247 | -104.039
7.5283 | -104.0151
7.532 | -103.9811
7.5356 | -103.9406
7.5392 | -103.9
7.5429 | -103.8505
7.5465 | -103.7991
7.5501 | -103.7486
7.5537 | -103.7023
7.5574 | -103.6666
7.561 | -103.6356
7.5646 | -103.6124
7.5683 | -103.5948
7.5719 | -103.5851
7.5755 | -103.5815
7.5791 | -103.5836
7.5828 | -103.5936
7.5864 | -103.609
7.59 | -103.6273
7.5937 | -103.6462
7.5973 | -103.6713
7.6009 | -103.6983
7.6045 | -103.7169
7.6082 | -103.7372
7.6118 | -103.7563
7.6154 | -103.7717
7.619 | -103.7881
7.6227 | -103.7954
7.6263 | -103.8005
7.6299 | -103.8
7.6336 | -103.7928
7.6372 | -103.7816
7.6408 | -103.7711
7.6444 | -103.7642
7.6481 | -103.7554
7.6517 | -103.7478
7.6553 | -103.7371
7.659 | -103.7264
7.6626 | -103.7221
7.6662 | -103.7181
7.6698 | -103.7139
7.6735 | -103.7101
7.6771 | -103.7039
7.6807 | -103.6958
7.6844 | -103.6864
7.688 | -103.6794
7.6916 | -103.6677
7.6952 | -103.6516
7.6989 | -103.6367
7.7025 | -103.6145
7.7061 | -103.5874
7.7098 | -103.5574
7.7134 | -103.5347
7.717 | -103.518
7.7206 | -103.5011
7.7243 | -103.4873
7.7279 | -103.4805
7.7315 | -103.4841
7.7351 | -103.4948
7.7388 | -103.5083
7.7424 | -103.5216
7.746 | -103.5388
7.7497 | -103.5576
7.7533 | -103.5778
7.7569 | -103.5946
7.7605 | -103.6134
7.7642 | -103.6237
7.7678 | -103.6293
7.7714 | -103.6319
7.7751 | -103.6267
7.7787 | -103.6192
7.7823 | -103.6114
7.7859 | -103.6013
7.7896 | -103.5905
7.7932 | -103.5793
7.7968 | -103.5658
7.8005 | -103.5525
7.8041 | -103.5402
7.8077 | -103.5312
7.8113 | -103.5204
7.815 | -103.515
7.8186 | -103.5156
7.8222 | -103.5179
7.8258 | -103.5137
7.8295 | -103.5155
7.8331 | -103.5286
7.8367 | -103.5435
7.8404 | -103.5596
7.844 | -103.5792
7.8476 | -103.601
7.8512 | -103.6425
7.8549 | -103.67
7.8585 | -103.7027
7.8621 | -103.7407
7.8658 | -103.7762
7.8694 | -103.8047
7.873 | -103.8349
7.8766 | -103.8599
7.8803 | -103.8815
7.8839 | -103.903
7.8875 | -103.9157
7.8912 | -103.9221
7.8948 | -103.9268
7.8984 | -103.9286
7.902 | -103.9303
7.9057 | -103.9286
7.9093 | -103.9225
7.9129 | -103.9243
7.9166 | -103.9269
7.9202 | -103.9254
7.9238 | -103.9234
7.9274 | -103.9244
7.9311 | -103.9246
7.9347 | -103.9209
7.9383 | -103.919
7.942 | -103.9201
7.9456 | -103.9106
7.9492 | -103.9021
7.9528 | -103.8938
7.9565 | -103.8841
7.9601 | -103.8743
7.9637 | -103.8684
7.9673 | -103.8615
7.971 | -103.8568
7.9746 | -103.8569
7.9782 | -103.8569
7.9819 | -103.855
7.9855 | -103.8648
7.9891 | -103.882
7.9927 | -103.8997
7.9964 | -103.9206
8 | -103.9388
8.0036 | -103.9558
8.0073 | -103.9751
8.0109 | -103.9892
8.0145 | -103.9972
8.0181 | -104.0051
8.0218 | -104.0115
8.0254 | -104.0167
8.029 | -104.0163
8.0327 | -104.0058
8.0363 | -103.995
8.0399 | -103.9877
8.0435 | -103.9753
8.0472 | -103.9607
8.0508 | -103.943
8.0544 | -103.9264
8.058 | -103.9075
8.0617 | -103.8924
8.0653 | -103.8861
8.0689 | -103.8766
8.0726 | -103.876
8.0762 | -103.8809
8.0798 | -103.8798
8.0834 | -103.8805
8.0871 | -103.8809
8.0907 | -103.8894
8.0943 | -103.8942
8.098 | -103.8956
8.1016 | -103.8937
8.1052 | -103.8891
8.1088 | -103.8847
8.1125 | -103.8842
8.1161 | -103.8933
8.1197 | -103.9085
8.1234 | -103.9249
8.127 | -103.9447
8.1306 | -103.9763
8.1342 | -104.0167
8.1379 | -104.0548
8.1415 | -104.0976
8.1451 | -104.1456
8.1488 | -104.1905
8.1524 | -104.2291
8.156 | -104.2699
8.1596 | -104.3036
8.1633 | -104.3233
8.1669 | -104.3422
8.1705 | -104.3538
8.1741 | -104.3497
8.1778 | -104.3368
8.1814 | -104.3147
8.185 | -104.2827
8.1887 | -104.2439
8.1923 | -104.2012
8.1959 | -104.1557
8.1995 | -104.1177
8.2032 | -104.0794
8.2068 | -104.0403
8.2104 | -104.0061
8.2141 | -103.9853
8.2177 | -103.9687
8.2213 | -103.9589
8.2249 | -103.9485
8.2286 | -103.9397
8.2322 | -103.932
8.2358 | -103.9258
8.2395 | -103.923
8.2431 | -103.9218
8.2467 | -103.9206
8.2503 | -103.9186
8.254 | -103.9138
8.2576 | -103.9136
8.2612 | -103.9119
8.2649 | -103.9095
8.2685 | -103.9106
8.2721 | -103.9151
8.2757 | -103.9193
8.2794 | -103.9283
8.283 | -103.942
8.2866 | -103.9519
8.2902 | -103.9669
8.2939 | -103.9882
8.2975 | -104.0124
8.3011 | -104.0436
8.3048 | -104.0765
8.3084 | -104.1112
8.312 | -104.1451
8.3156 | -104.1758
8.3193 | -104.2035
8.3229 | -104.2318
8.3265 | -104.2608
8.3302 | -104.2854
8.3338 | -104.3119
8.3374 | -104.3387
8.341 | -104.3706
8.3447 | -104.4027
8.3483 | -104.4341
8.3519 | -104.4569
8.3556 | -104.4702
8.3592 | -104.4804
8.3628 | -104.4826
8.3664 | -104.482
8.3701 | -104.4759
8.3737 | -104.4435
8.3773 | -104.4215
8.381 | -104.3926
8.3846 | -104.3598
8.3882 | -104.3246
8.3918 | -104.2902
8.3955 | -104.2527
8.3991 | -104.2137
8.4027 | -104.1806
8.4063 | -104.1529
8.41 | -104.1258
8.4136 | -104.0969
8.4172 | -104.0708
8.4209 | -104.0469
8.4245 | -104.0233
8.4281 | -104.0058
8.4317 | -103.9881
8.4354 | -103.9722
8.439 | -103.963
8.4426 | -103.9478
8.4463 | -103.9359
8.4499 | -103.9273
8.4535 | -103.9163
8.4571 | -103.9053
8.4608 | -103.8966
8.4644 | -103.8795
8.468 | -103.8639
8.4717 | -103.8451
8.4753 | -103.8272
8.4789 | -103.8065
8.4825 | -103.7815
8.4862 | -103.7564
8.4898 | -103.7284
8.4934 | -103.7057
8.4971 | -103.6901
8.5007 | -103.679
8.5043 | -103.6744
8.5079 | -103.675
8.5116 | -103.6842
8.5152 | -103.6985
8.5188 | -103.7143
8.5224 | -103.7358
8.5261 | -103.7574
8.5297 | -103.7803
8.5333 | -103.8049
8.537 | -103.8307
8.5406 | -103.8579
8.5442 | -103.8795
8.5478 | -103.9009
8.5515 | -103.922
8.5551 | -103.9416
8.5587 | -103.9621
8.5624 | -103.9824
8.566 | -104.0046
8.5696 | -104.032
8.5732 | -104.0529
8.5769 | -104.0789
8.5805 | -104.1069
8.5841 | -104.1343
8.5878 | -104.1649
8.5914 | -104.2028
8.595 | -104.2433
8.5986 | -104.2821
8.6023 | -104.3191
8.6059 | -104.3512
8.6095 | -104.3771
8.6132 | -104.3978
8.6168 | -104.415
8.6204 | -104.4227
8.624 | -104.4267
8.6277 | -104.4295
8.6313 | -104.4245
8.6349 | -104.4152
8.6385 | -104.3988
8.6422 | -104.383
8.6458 | -104.3688
8.6494 | -104.3528
8.6531 | -104.3392
8.6567 | -104.3252
8.6603 | -104.3166
8.6639 | -104.3095
8.6676 | -104.3057
8.6712 | -104.3037
8.6748 | -104.3013
8.6785 | -104.3025
8.6821 | -104.303
8.6857 | -104.3044
8.6893 | -104.3092
8.693 | -104.3139
8.6966 | -104.3168
8.7002 | -104.3177
8.7039 | -104.3136
8.7075 | -104.3073
8.7111 | -104.299
8.7147 | -104.2857
8.7184 | -104.2644
8.722 | -104.2356
8.7256 | -104.2074
8.7293 | -104.1739
8.7329 | -104.1398
8.7365 | -104.1061
8.7401 | -104.0699
8.7438 | -104.0308
8.7474 | -103.9956
8.751 | -103.9643
8.7546 | -103.9347
8.7583 | -103.9149
8.7619 | -103.8983
8.7655 | -103.885
8.7692 | -103.8685
8.7728 | -103.8588
8.7764 | -103.8569
8.78 | -103.8561
8.7837 | -103.8601
8.7873 | -103.8643
8.7909 | -103.8669
8.7946 | -103.8683
8.7982 | -103.8714
8.8018 | -103.8752
8.8054 | -103.8806
8.8091 | -103.8865
8.8127 | -103.8927
8.8163 | -103.8975
8.82 | -103.9047
8.8236 | -103.917
8.8272 | -103.9341
8.8308 | -103.9475
8.8345 | -103.9626
8.8381 | -103.9743
8.8417 | -103.9853
8.8454 | -103.9963
8.849 | -104.0096
8.8526 | -104.016
8.8562 | -104.0148
8.8599 | -104.0145
8.8635 | -104.0127
8.8671 | -104.0089
8.8707 | -104.0046
8.8744 | -103.9997
8.878 | -103.9923
8.8816 | -103.9812
8.8853 | -103.9672
8.8889 | -103.9553
8.8925 | -103.9459
8.8961 | -103.9306
8.8998 | -103.9207
8.9034 | -103.9086
8.907 | -103.8903
8.9107 | -103.8725
8.9143 | -103.8572
8.9179 | -103.8444
8.9215 | -103.8347
8.9252 | -103.826
8.9288 | -103.8171
8.9324 | -103.807
8.9361 | -103.7986
8.9397 | -103.7956
8.9433 | -103.7963
8.9469 | -103.8013
8.9506 | -103.8037
8.9542 | -103.8086
8.9578 | -103.8205
8.9615 | -103.837
8.9651 | -103.8526
8.9687 | -103.8735
8.9723 | -103.8931
8.976 | -103.9096
8.9796 | -103.9234
8.9832 | -103.9404
8.9868 | -103.9602
8.9905 | -103.9704
8.9941 | -103.9784
8.9977 | -103.9891
9.0014 | -104.004
9.005 | -104.0189
9.0086 | -104.0341
9.0122 | -104.0534
9.0159 | -104.0752
9.0195 | -104.0996
9.0231 | -104.1256
9.0268 | -104.1545
9.0304 | -104.1945
9.034 | -104.2361
9.0376 | -104.2756
9.0413 | -104.3214
9.0449 | -104.3712
9.0485 | -104.4199
9.0522 | -104.4728
9.0558 | -104.5338
9.0594 | -104.5933
9.063 | -104.6525
9.0667 | -104.7122
9.0703 | -104.7639
9.0739 | -104.8143
9.0776 | -104.8585
9.0812 | -104.8927
9.0848 | -104.9232
9.0884 | -104.9494
9.0921 | -104.9595
9.0957 | -104.963
9.0993 | -104.9651
9.1029 | -104.9598
9.1066 | -104.9518
9.1102 | -104.9409
9.1138 | -104.9291
9.1175 | -104.9139
9.1211 | -104.8971
9.1247 | -104.881
9.1283 | -104.8661
9.132 | -104.8536
9.1356 | -104.8381
9.1392 | -104.8274
9.1429 | -104.8196
9.1465 | -104.8185
9.1501 | -104.8183
9.1537 | -104.8179
9.1574 | -104.8179
9.161 | -104.817
9.1646 | -104.8116
9.1683 | -104.8076
9.1719 | -104.8062
9.1755 | -104.802
9.1791 | -104.7957
9.1828 | -104.7896
9.1864 | -104.7813
9.19 | -104.7707
9.1937 | -104.7621
9.1973 | -104.7524
9.2009 | -104.7434
9.2045 | -104.7396
9.2082 | -104.734
9.2118 | -104.7299
9.2154 | -104.7304
9.219 | -104.733
9.2227 | -104.7338
9.2263 | -104.7392
9.2299 | -104.744
9.2336 | -104.748
9.2372 | -104.7481
9.2408 | -104.7483
9.2444 | -104.7508
9.2481 | -104.7544
9.2517 | -104.7616
9.2553 | -104.7725
9.259 | -104.787
9.2626 | -104.8036
9.2662 | -104.8232
9.2698 | -104.8455
9.2735 | -104.8715
9.2771 | -104.9012
9.2807 | -104.921
9.2844 | -104.9402
9.288 | -104.9564
9.2916 | -104.9718
9.2952 | -104.9834
9.2989 | -104.9913
9.3025 | -104.9965
9.3061 | -104.9956
9.3098 | -104.9913
9.3134 | -104.9828
9.317 | -104.9693
9.3206 | -104.9552
9.3243 | -104.9394
9.3279 | -104.9271
9.3315 | -104.9165
9.3351 | -104.9042
9.3388 | -104.8895
9.3424 | -104.8739
9.346 | -104.8634
9.3497 | -104.8556
9.3533 | -104.8475
9.3569 | -104.8435
9.3605 | -104.8426
9.3642 | -104.8439
9.3678 | -104.8461
9.3714 | -104.8513
9.3751 | -104.8566
9.3787 | -104.8596
9.3823 | -104.8629
9.3859 | -104.8635
9.3896 | -104.8616
9.3932 | -104.8569
9.3968 | -104.8489
9.4005 | -104.8353
9.4041 | -104.8201
9.4077 | -104.8042
9.4113 | -104.79
9.415 | -104.7782
9.4186 | -104.7613
9.4222 | -104.7459
9.4259 | -104.7314
9.4295 | -104.7151
9.4331 | -104.6995
9.4367 | -104.6833
9.4404 | -104.6686
9.444 | -104.6568
9.4476 | -104.6477
9.4512 | -104.6452
9.4549 | -104.643
9.4585 | -104.6358
9.4621 | -104.6346
9.4658 | -104.6363
9.4694 | -104.6396
9.473 | -104.6485
9.4766 | -104.6618
9.4803 | -104.6783
9.4839 | -104.696
9.4875 | -104.7154
9.4912 | -104.7403
9.4948 | -104.765
9.4984 | -104.7915
9.502 | -104.8192
9.5057 | -104.8459
9.5093 | -104.8718
9.5129 | -104.898
9.5166 | -104.9225
9.5202 | -104.9406
9.5238 | -104.9554
9.5274 | -104.9695
9.5311 | -104.9832
9.5347 | -104.9894
9.5383 | -104.9938
9.5419 | -104.9929
9.5456 | -104.9801
9.5492 | -104.9571
9.5528 | -104.9356
9.5565 | -104.9122
9.5601 | -104.8901
9.5637 | -104.8672
9.5673 | -104.8403
9.571 | -104.8156
9.5746 | -104.7945
9.5782 | -104.777
9.5819 | -104.7615
9.5855 | -104.751
9.5891 | -104.743
9.5927 | -104.7361
9.5964 | -104.7387
9.6 | -104.751
9.6036 | -104.7631
9.6073 | -104.774
9.6109 | -104.7905
9.6145 | -104.8083
9.6181 | -104.8312
9.6218 | -104.8573
9.6254 | -104.8799
9.629 | -104.9004
9.6327 | -104.9264
9.6363 | -104.9554
9.6399 | -104.9856
9.6435 | -105.0167
9.6472 | -105.0469
9.6508 | -105.0696
9.6544 | -105.0866
9.658 | -105.0976
9.6617 | -105.1016
9.6653 | -105.1049
9.6689 | -105.1013
9.6726 | -105.0922
9.6762 | -105.0824
9.6798 | -105.0666
9.6834 | -105.047
9.6871 | -105.03
9.6907 | -105.0172
9.6943 | -104.9996
9.698 | -104.9796
9.7016 | -104.9598
9.7052 | -104.945
9.7088 | -104.9379
9.7125 | -104.9308
9.7161 | -104.9224
9.7197 | -104.9158
9.7234 | -104.915
9.727 | -104.9133
9.7306 | -104.9154
9.7342 | -104.9169
9.7379 | -104.9174
9.7415 | -104.92
9.7451 | -104.9242
9.7488 | -104.9306
9.7524 | -104.9331
9.756 | -104.9324
9.7596 | -104.9287
9.7633 | -104.9278
9.7669 | -104.9308
9.7705 | -104.9343
9.7742 | -104.9313
9.7778 | -104.924
9.7814 | -104.9142
9.785 | -104.9035
9.7887 | -104.8926
9.7923 | -104.8848
9.7959 | -104.8742
9.7995 | -104.8631
9.8032 | -104.8494
9.8068 | -104.83
9.8104 | -104.8096
9.8141 | -104.7885
9.8177 | -104.768
9.8213 | -104.7482
9.8249 | -104.7291
9.8286 | -104.7083
9.8322 | -104.6881
9.8358 | -104.6695
9.8395 | -104.6505
9.8431 | -104.6323
9.8467 | -104.617
9.8503 | -104.6045
9.854 | -104.5973
9.8576 | -104.5947
9.8612 | -104.5934
9.8649 | -104.5884
9.8685 | -104.5875
9.8721 | -104.588
9.8757 | -104.5946
9.8794 | -104.6038
9.883 | -104.6097
9.8866 | -104.6172
9.8902 | -104.6282
9.8939 | -104.643
9.8975 | -104.6582
9.9011 | -104.6704
9.9048 | -104.6826
9.9084 | -104.6941
9.912 | -104.7049
9.9156 | -104.7176
9.9193 | -104.7336
9.9229 | -104.752
9.9265 | -104.7699
9.9302 | -104.7844
9.9338 | -104.797
9.9374 | -104.8095
9.941 | -104.8196
9.9447 | -104.8198
9.9483 | -104.8316
9.9519 | -104.8381
9.9556 | -104.8391
9.9592 | -104.8334
9.9628 | -104.8243
9.9664 | -104.8133
9.9701 | -104.8035
9.9737 | -104.7917
9.9773 | -104.7776
9.981 | -104.7591
9.9846 | -104.7466
9.9882 | -104.7368
9.9918 | -104.7238
9.9955 | -104.7081
9.9991 | -104.6945
10.0027 | -104.6829
10.0063 | -104.6741
10.01 | -104.6688
10.0136 | -104.6671
10.0172 | -104.668
10.0209 | -104.6711
10.0245 | -104.6746
10.0281 | -104.6831
10.0317 | -104.6937
10.0354 | -104.7069
10.039 | -104.7213
10.0426 | -104.7423
10.0463 | -104.7663
10.0499 | -104.7901
10.0535 | -104.8155
10.0571 | -104.8441
10.0608 | -104.8717
10.0644 | -104.8969
10.068 | -104.9195
10.0717 | -104.9456
10.0753 | -104.97
10.0789 | -104.9914
10.0825 | -105.0118
10.0862 | -105.0305
10.0898 | -105.0512
10.0934 | -105.0652
10.0971 | -105.0747
10.1007 | -105.0873
10.1043 | -105.1001
10.1079 | -105.114
10.1116 | -105.1269
10.1152 | -105.1347
10.1188 | -105.1365
10.1224 | -105.1383
10.1261 | -105.1396
10.1297 | -105.1393
10.1333 | -105.1345
10.137 | -105.1241
10.1406 | -105.1089
10.1442 | -105.0932
10.1478 | -105.078
10.1515 | -105.0605
10.1551 | -105.04
10.1587 | -105.0185
10.1624 | -104.9885
10.166 | -104.9569
10.1696 | -104.9261
10.1732 | -104.8998
10.1769 | -104.8746
10.1805 | -104.8522
10.1841 | -104.828
10.1878 | -104.8086
10.1914 | -104.7877
10.195 | -104.7707
10.1986 | -104.7544
10.2023 | -104.7413
10.2059 | -104.7337
10.2095 | -104.7294
10.2132 | -104.731
10.2168 | -104.7319
10.2204 | -104.7308
10.224 | -104.7346
10.2277 | -104.7375
10.2313 | -104.7408
10.2349 | -104.7458
10.2385 | -104.7506
10.2422 | -104.758
10.2458 | -104.7663
10.2494 | -104.7743
10.2531 | -104.7821
10.2567 | -104.7892
10.2603 | -104.7963
10.2639 | -104.805
10.2676 | -104.8147
10.2712 | -104.8194
10.2748 | -104.8244
10.2785 | -104.836
10.2821 | -104.8462
10.2857 | -104.8594
10.2893 | -104.8765
10.293 | -104.8903
10.2966 | -104.8996
10.3002 | -104.9134
10.3039 | -104.9289
10.3075 | -104.9419
10.3111 | -104.9602
10.3147 | -104.9744
10.3184 | -104.9833
10.322 | -104.9877
10.3256 | -104.9908
10.3293 | -104.9907
10.3329 | -104.9937
10.3365 | -104.9944
10.3401 | -104.9909
10.3438 | -104.9857
10.3474 | -104.9792
10.351 | -104.972
10.3546 | -104.9611
10.3583 | -104.9502
10.3619 | -104.937
10.3655 | -104.9314
10.3692 | -104.9276
10.3728 | -104.9213
10.3764 | -104.9196
10.38 | -104.9173
10.3837 | -104.9138
10.3873 | -104.9086
10.3909 | -104.9051
10.3946 | -104.9059
10.3982 | -104.9103
10.4018 | -104.9169
10.4054 | -104.926
10.4091 | -104.9335
10.4127 | -104.9396
10.4163 | -104.9502
10.42 | -104.9682
10.4236 | -104.9861
10.4272 | -104.9997
10.4308 | -105.015
10.4345 | -105.031
10.4381 | -105.045
10.4417 | -105.0611
10.4454 | -105.0735
10.449 | -105.0821
10.4526 | -105.0923
10.4562 | -105.1006
10.4599 | -105.1065
10.4635 | -105.1173
10.4671 | -105.1215
10.4707 | -105.1304
10.4744 | -105.1376
10.478 | -105.1455
10.4816 | -105.1532
10.4853 | -105.1593
10.4889 | -105.1669
10.4925 | -105.1787
10.4961 | -105.1877
10.4998 | -105.1975
10.5034 | -105.208
10.507 | -105.2187
10.5107 | -105.2303
10.5143 | -105.2441
10.5179 | -105.2546
10.5215 | -105.2629
10.5252 | -105.2715
10.5288 | -105.2795
10.5324 | -105.2865
10.5361 | -105.2953
10.5397 | -105.3049
10.5433 | -105.3094
10.5469 | -105.3089
10.5506 | -105.3035
10.5542 | -105.2975
10.5578 | -105.2883
10.5615 | -105.2795
10.5651 | -105.2696
10.5687 | -105.2525
10.5723 | -105.2333
10.576 | -105.2137
10.5796 | -105.1976
10.5832 | -105.1816
10.5868 | -105.1643
10.5905 | -105.15
10.5941 | -105.1365
10.5977 | -105.1242
10.6014 | -105.1133
10.605 | -105.1038
10.6086 | -105.0984
10.6122 | -105.0955
10.6159 | -105.0902
10.6195 | -105.0875
10.6231 | -105.0855
10.6268 | -105.0813
10.6304 | -105.0786
10.634 | -105.0711
10.6376 | -105.0643
10.6413 | -105.0595
10.6449 | -105.0501
10.6485 | -105.0366
10.6522 | -105.0265
10.6558 | -105.0149
10.6594 | -105.0028
10.663 | -104.9937
10.6667 | -104.9847
10.6703 | -104.9751
10.6739 | -104.9639
10.6776 | -104.9533
10.6812 | -104.9439
10.6848 | -104.9328
10.6884 | -104.9248
10.6921 | -104.917
10.6957 | -104.9114
10.6993 | -104.9096
10.7029 | -104.9069
10.7066 | -104.9055
10.7102 | -104.906
10.7138 | -104.9072
10.7175 | -104.9084
10.7211 | -104.9107
10.7247 | -104.9135
10.7283 | -104.9127
10.732 | -104.9127
10.7356 | -104.9127
10.7392 | -104.9143
10.7429 | -104.9165
10.7465 | -104.9188
10.7501 | -104.9205
10.7537 | -104.9188
10.7574 | -104.9156
10.761 | -104.9109
10.7646 | -104.908
10.7683 | -104.9035
10.7719 | -104.9031
10.7755 | -104.9055
10.7791 | -104.9074
10.7828 | -104.9084
10.7864 | -104.9095
10.79 | -104.9134
10.7937 | -104.9168
10.7973 | -104.9174
10.8009 | -104.9231
10.8045 | -104.9335
10.8082 | -104.94
10.8118 | -104.9455
10.8154 | -104.9487
10.819 | -104.9528
10.8227 | -104.9579
10.8263 | -104.9647
10.8299 | -104.9761
10.8336 | -104.9874
10.8372 | -104.9981
10.8408 | -105.0049
10.8444 | -105.0088
10.8481 | -105.014
10.8517 | -105.016
10.8553 | -105.017
10.859 | -105.0193
10.8626 | -105.0177
10.8662 | -105.0118
10.8698 | -105.0051
10.8735 | -104.9963
10.8771 | -104.9858
10.8807 | -104.9745
10.8844 | -104.9628
10.888 | -104.9511
10.8916 | -104.9363
10.8952 | -104.9217
10.8989 | -104.908
10.9025 | -104.8981
10.9061 | -104.8943
10.9098 | -104.8955
10.9134 | -104.8989
10.917 | -104.9058
10.9206 | -104.9162
10.9243 | -104.9322
10.9279 | -104.9497
10.9315 | -104.972
10.9351 | -105.003
10.9388 | -105.0411
10.9424 | -105.0837
10.946 | -105.1282
10.9497 | -105.1737
10.9533 | -105.2221
10.9569 | -105.269
10.9605 | -105.317
10.9642 | -105.3657
10.9678 | -105.4132
10.9714 | -105.4613
10.9751 | -105.5028
10.9787 | -105.5402
10.9823 | -105.5703
10.9859 | -105.5948
10.9896 | -105.5918
10.9932 | -105.608
10.9968 | -105.6175
11.0005 | -105.6257
11.0041 | -105.6256
11.0077 | -105.6192
11.0113 | -105.6128
11.015 | -105.6029
11.0186 | -105.5862
11.0222 | -105.5666
11.0259 | -105.5447
11.0295 | -105.5174
11.0331 | -105.4886
11.0367 | -105.4629
11.0404 | -105.4325
11.044 | -105.4016
11.0476 | -105.3724
11.0512 | -105.3455
11.0549 | -105.3176
11.0585 | -105.2911
11.0621 | -105.2632
11.0658 | -105.2385
11.0694 | -105.2181
11.073 | -105.1985
11.0766 | -105.1803
11.0803 | -105.1637
11.0839 | -105.1518
11.0875 | -105.142
11.0912 | -105.1348
11.0948 | -105.1275
11.0984 | -105.1224
11.102 | -105.1214
11.1057 | -105.122
11.1093 | -105.1259
11.1129 | -105.1332
11.1166 | -105.1462
11.1202 | -105.1621
11.1238 | -105.1793
11.1274 | -105.1962
11.1311 | -105.2122
11.1347 | -105.2329
11.1383 | -105.2542
11.1419 | -105.2785
11.1456 | -105.3001
11.1492 | -105.3185
11.1528 | -105.3385
11.1565 | -105.3541
11.1601 | -105.3665
11.1637 | -105.3738
11.1673 | -105.3835
11.171 | -105.3893
11.1746 | -105.3906
11.1782 | -105.3879
11.1819 | -105.3868
11.1855 | -105.3881
11.1891 | -105.3892
11.1927 | -105.3883
11.1964 | -105.3855
11.2 | -105.3792
11.2036 | -105.3757
11.2073 | -105.3704
11.2109 | -105.3635
11.2145 | -105.3584
11.2181 | -105.3542
11.2218 | -105.3516
11.2254 | -105.3464
11.229 | -105.3363
11.2327 | -105.3273
11.2363 | -105.3207
11.2399 | -105.3179
11.2435 | -105.318
11.2472 | -105.3145
11.2508 | -105.313
11.2544 | -105.3117
11.2581 | -105.3107
11.2617 | -105.3107
11.2653 | -105.3107
11.2689 | -105.3147
11.2726 | -105.3201
11.2762 | -105.325
11.2798 | -105.3287
11.2834 | -105.3326
11.2871 | -105.3375
11.2907 | -105.3457
11.2943 | -105.3555
11.298 | -105.3605
11.3016 | -105.3648
11.3052 | -105.3683
11.3088 | -105.3726
11.3125 | -105.3779
11.3161 | -105.3785
11.3197 | -105.38
11.3234 | -105.3801
11.327 | -105.3782
11.3306 | -105.3723
11.3342 | -105.3668
11.3379 | -105.3576
11.3415 | -105.3487
11.3451 | -105.3408
11.3488 | -105.3273
11.3524 | -105.3089
11.356 | -105.2951
11.3596 | -105.2804
11.3633 | -105.2641
11.3669 | -105.249
11.3705 | -105.2334
11.3742 | -105.2193
11.3778 | -105.2067
11.3814 | -105.196
11.385 | -105.1847
11.3887 | -105.1769
11.3923 | -105.1719
11.3959 | -105.1672
11.3995 | -105.1661
11.4032 | -105.1661
11.4068 | -105.1672
11.4104 | -105.1732
11.4141 | -105.1791
11.4177 | -105.1838
11.4213 | -105.1949
11.4249 | -105.2079
11.4286 | -105.2192
11.4322 | -105.2295
11.4358 | -105.2399
11.4395 | -105.2528
11.4431 | -105.2644
11.4467 | -105.275
11.4503 | -105.2854
11.454 | -105.2957
11.4576 | -105.3008
11.4612 | -105.3077
11.4649 | -105.3162
11.4685 | -105.3226
11.4721 | -105.329
11.4757 | -105.3399
11.4794 | -105.3519
11.483 | -105.3625
11.4866 | -105.3787
11.4902 | -105.3938
11.4939 | -105.4083
11.4975 | -105.4236
11.5011 | -105.4403
11.5048 | -105.4574
11.5084 | -105.4704
11.512 | -105.4823
11.5156 | -105.4946
11.5193 | -105.5055
11.5229 | -105.5135
11.5265 | -105.5198
11.5302 | -105.5222
11.5338 | -105.5227
11.5374 | -105.5235
11.541 | -105.5177
11.5447 | -105.5073
11.5483 | -105.4959
11.5519 | -105.488
11.5556 | -105.4803
11.5592 | -105.4704
11.5628 | -105.4581
11.5664 | -105.4416
11.5701 | -105.4263
11.5737 | -105.4136
11.5773 | -105.4003
11.581 | -105.3875
11.5846 | -105.3786
11.5882 | -105.3678
11.5918 | -105.3566
11.5955 | -105.3469
11.5991 | -105.3384
11.6027 | -105.3306
11.6063 | -105.3225
11.61 | -105.3173
11.6136 | -105.3108
11.6172 | -105.3097
11.6209 | -105.3081
11.6245 | -105.309
11.6281 | -105.3103
11.6317 | -105.309
11.6354 | -105.3123
11.639 | -105.3159
11.6426 | -105.3193
11.6463 | -105.328
11.6499 | -105.3373
11.6535 | -105.3458
11.6571 | -105.3583
11.6608 | -105.3708
11.6644 | -105.38
11.668 | -105.3912
11.6717 | -105.4017
11.6753 | -105.4099
11.6789 | -105.4143
11.6825 | -105.4162
11.6862 | -105.42
11.6898 | -105.4228
11.6934 | -105.4175
11.6971 | -105.4102
11.7007 | -105.4042
11.7043 | -105.3929
11.7079 | -105.3817
11.7116 | -105.3686
11.7152 | -105.3527
11.7188 | -105.3367
11.7224 | -105.3185
11.7261 | -105.3008
11.7297 | -105.2828
11.7333 | -105.2673
11.737 | -105.2534
11.7406 | -105.2399
11.7442 | -105.2254
11.7478 | -105.2124
11.7515 | -105.2042
11.7551 | -105.1963
11.7587 | -105.1871
11.7624 | -105.18
11.766 | -105.1764
11.7696 | -105.1765
11.7732 | -105.1756
11.7769 | -105.1765
11.7805 | -105.1777
11.7841 | -105.1824
11.7878 | -105.1849
11.7914 | -105.1879
11.795 | -105.1916
11.7986 | -105.1948
11.8023 | -105.1985
11.8059 | -105.2024
11.8095 | -105.2073
11.8132 | -105.2099
11.8168 | -105.2122
11.8204 | -105.2123
11.824 | -105.2113
11.8277 | -105.2108
11.8313 | -105.2107
11.8349 | -105.2109
11.8385 | -105.208
11.8422 | -105.2066
11.8458 | -105.2042
11.8494 | -105.2065
11.8531 | -105.2057
11.8567 | -105.2032
11.8603 | -105.2032
11.8639 | -105.2066
11.8676 | -105.2072
11.8712 | -105.2056
11.8748 | -105.2063
11.8785 | -105.2064
11.8821 | -105.2089
11.8857 | -105.2102
11.8893 | -105.2138
11.893 | -105.2184
11.8966 | -105.2238
11.9002 | -105.2305
11.9039 | -105.2354
11.9075 | -105.2421
11.9111 | -105.2497
11.9147 | -105.2592
11.9184 | -105.2706
11.922 | -105.2833
11.9256 | -105.2943
11.9293 | -105.3066
11.9329 | -105.3237
11.9365 | -105.3399
11.9401 | -105.3545
11.9438 | -105.3712
11.9474 | -105.3879
11.951 | -105.4069
11.9546 | -105.4212
11.9583 | -105.4377
11.9619 | -105.457
11.9655 | -105.4777
11.9692 | -105.4987
11.9728 | -105.5175
11.9764 | -105.5322
11.98 | -105.5477
11.9837 | -105.5632
11.9873 | -105.5782
11.9909 | -105.5971
11.9946 | -105.6157
11.9982 | -105.6317
12.0018 | -105.6472
12.0054 | -105.6633
12.0091 | -105.6786
12.0127 | -105.6891
12.0163 | -105.7021
12.02 | -105.7096
12.0236 | -105.7171
12.0272 | -105.7165
12.0308 | -105.7163
12.0345 | -105.7158
12.0381 | -105.7045
12.0417 | -105.6991
12.0454 | -105.6906
12.049 | -105.6827
12.0526 | -105.6729
12.0562 | -105.6603
12.0599 | -105.6485
12.0635 | -105.6316
12.0671 | -105.6135
12.0707 | -105.5966
12.0744 | -105.58
12.078 | -105.5634
12.0816 | -105.5454
12.0853 | -105.5267
12.0889 | -105.5087
12.0925 | -105.491
12.0961 | -105.4695
12.0998 | -105.4503
12.1034 | -105.434
12.107 | -105.4172
12.1107 | -105.4031
12.1143 | -105.3893
12.1179 | -105.3786
12.1215 | -105.3706
12.1252 | -105.3632
12.1288 | -105.3586
12.1324 | -105.3523
12.1361 | -105.347
12.1397 | -105.3429
12.1433 | -105.3386
12.1469 | -105.3397
12.1506 | -105.3428
12.1542 | -105.3458
12.1578 | -105.3515
12.1615 | -105.3549
12.1651 | -105.3567
12.1687 | -105.3617
12.1723 | -105.3657
12.176 | -105.3703
12.1796 | -105.3762
12.1832 | -105.3842
12.1868 | -105.3948
12.1905 | -105.4053
12.1941 | -105.416
12.1977 | -105.4289
12.2014 | -105.4476
12.205 | -105.4614
12.2086 | -105.4738
12.2122 | -105.4903
12.2159 | -105.5079
12.2195 | -105.528
12.2231 | -105.5436
12.2268 | -105.5585
12.2304 | -105.5714
12.234 | -105.5818
12.2376 | -105.5927
12.2413 | -105.6034
12.2449 | -105.6134
12.2485 | -105.62
12.2522 | -105.6281
12.2558 | -105.6347
12.2594 | -105.6386
12.263 | -105.6411
12.2667 | -105.6428
12.2703 | -105.64
12.2739 | -105.6348
12.2776 | -105.6297
12.2812 | -105.6209
12.2848 | -105.6095
12.2884 | -105.5961
12.2921 | -105.5835
12.2957 | -105.5682
12.2993 | -105.5494
12.3029 | -105.5312
12.3066 | -105.5176
12.3102 | -105.5046
12.3138 | -105.491
12.3175 | -105.4769
12.3211 | -105.4652
12.3247 | -105.455
12.3283 | -105.4414
12.332 | -105.4289
12.3356 | -105.4206
12.3392 | -105.4177
12.3429 | -105.4135
12.3465 | -105.408
12.3501 | -105.4073
12.3537 | -105.4078
12.3574 | -105.4064
12.361 | -105.4046
12.3646 | -105.4011
12.3683 | -105.3976
12.3719 | -105.3976
12.3755 | -105.3975
12.3791 | -105.3945
12.3828 | -105.3924
12.3864 | -105.3946
12.39 | -105.3946
12.3937 | -105.395
12.3973 | -105.3971
12.4009 | -105.4053
12.4045 | -105.4124
12.4082 | -105.4219
12.4118 | -105.437
12.4154 | -105.453
12.419 | -105.4694
12.4227 | -105.4853
12.4263 | -105.5004
12.4299 | -105.5161
12.4336 | -105.5356
12.4372 | -105.5554
12.4408 | -105.575
12.4444 | -105.5929
12.4481 | -105.6119
12.4517 | -105.6344
12.4553 | -105.6541
12.459 | -105.671
12.4626 | -105.685
12.4662 | -105.6993
12.4698 | -105.7132
12.4735 | -105.7239
12.4771 | -105.7335
12.4807 | -105.7389
12.4844 | -105.7374
12.488 | -105.7345
12.4916 | -105.7297
12.4952 | -105.7192
12.4989 | -105.7087
12.5025 | -105.6937
12.5061 | -105.677
12.5098 | -105.6576
12.5134 | -105.6322
12.517 | -105.6066
12.5206 | -105.5854
12.5243 | -105.5605
12.5279 | -105.5353
12.5315 | -105.5077
12.5351 | -105.4819
12.5388 | -105.4595
12.5424 | -105.4417
12.546 | -105.4269
12.5497 | -105.4136
12.5533 | -105.4018
12.5569 | -105.395
12.5605 | -105.4009
12.5642 | -105.3978
12.5678 | -105.3946
12.5714 | -105.398
12.5751 | -105.4007
12.5787 | -105.4065
12.5823 | -105.4138
12.5859 | -105.4221
12.5896 | -105.43
12.5932 | -105.4398
12.5968 | -105.4506
12.6005 | -105.4602
12.6041 | -105.4734
12.6077 | -105.4877
12.6113 | -105.4992
12.615 | -105.5079
12.6186 | -105.515
12.6222 | -105.5185
12.6258 | -105.5231
12.6295 | -105.5288
12.6331 | -105.5281
12.6367 | -105.5229
12.6404 | -105.5171
12.644 | -105.5089
12.6476 | -105.4967
12.6512 | -105.4823
12.6549 | -105.4693
12.6585 | -105.4539
12.6621 | -105.4368
12.6658 | -105.4219
12.6694 | -105.4041
12.673 | -105.3873
12.6766 | -105.3748
12.6803 | -105.3614
12.6839 | -105.3495
12.6875 | -105.3438
12.6912 | -105.3404
12.6948 | -105.3368
12.6984 | -105.3337
12.702 | -105.3346
12.7057 | -105.3396
12.7093 | -105.3421
12.7129 | -105.3508
12.7166 | -105.3648
12.7202 | -105.3814
12.7238 | -105.3987
12.7274 | -105.4164
12.7311 | -105.4383
12.7347 | -105.4622
12.7383 | -105.4844
12.7419 | -105.5074
12.7456 | -105.5283
12.7492 | -105.5509
12.7528 | -105.5714
12.7565 | -105.5884
12.7601 | -105.6075
12.7637 | -105.6235
12.7673 | -105.6378
12.771 | -105.6489
12.7746 | -105.6592
12.7782 | -105.6642
12.7819 | -105.6664
12.7855 | -105.6674
12.7891 | -105.6653
12.7927 | -105.664
12.7964 | -105.6611
12.8 | -105.6535
12.8036 | -105.642
12.8073 | -105.6337
12.8109 | -105.6198
12.8145 | -105.6024
12.8181 | -105.5901
12.8218 | -105.5786
12.8254 | -105.5634
12.829 | -105.551
12.8327 | -105.5382
12.8363 | -105.5234
12.8399 | -105.5095
12.8435 | -105.4965
12.8472 | -105.4838
12.8508 | -105.4716
12.8544 | -105.462
12.8581 | -105.4539
12.8617 | -105.4481
12.8653 | -105.4444
12.8689 | -105.4413
12.8726 | -105.4379
12.8762 | -105.4374
12.8798 | -105.4415
12.8834 | -105.4464
12.8871 | -105.4522
12.8907 | -105.4616
12.8943 | -105.4743
12.898 | -105.4877
12.9016 | -105.4982
12.9052 | -105.5074
12.9088 | -105.5188
12.9125 | -105.5329
12.9161 | -105.5436
12.9197 | -105.5571
12.9234 | -105.576
12.927 | -105.5938
12.9306 | -105.6124
12.9342 | -105.6266
12.9379 | -105.6409
12.9415 | -105.6575
12.9451 | -105.6744
12.9488 | -105.6912
12.9524 | -105.7079
12.956 | -105.724
12.9596 | -105.7396
12.9633 | -105.7534
12.9669 | -105.7635
12.9705 | -105.7716
12.9741 | -105.7779
12.9778 | -105.7793
12.9814 | -105.7779
12.985 | -105.7783
12.9887 | -105.7779
12.9923 | -105.7755
12.9959 | -105.7709
12.9995 | -105.7624
13.0032 | -105.75
13.0068 | -105.7349
13.0104 | -105.7197
13.0141 | -105.6993
13.0177 | -105.679
13.0213 | -105.6592
13.0249 | -105.6354
13.0286 | -105.6116
13.0322 | -105.5897
13.0358 | -105.5633
13.0395 | -105.5343
13.0431 | -105.5069
13.0467 | -105.4802
13.0503 | -105.4531
13.054 | -105.4291
13.0576 | -105.4068
13.0612 | -105.387
13.0649 | -105.3706
13.0685 | -105.355
13.0721 | -105.3387
13.0757 | -105.3269
13.0794 | -105.3193
13.083 | -105.3288
13.0866 | -105.3264
13.0902 | -105.3277
13.0939 | -105.3311
13.0975 | -105.3349
13.1011 | -105.3409
13.1048 | -105.3512
13.1084 | -105.3615
13.112 | -105.3681
13.1156 | -105.3774
13.1193 | -105.3892
13.1229 | -105.4011
13.1265 | -105.4144
13.1302 | -105.4253
13.1338 | -105.4345
13.1374 | -105.4411
13.141 | -105.4483
13.1447 | -105.4558
13.1483 | -105.4593
13.1519 | -105.4573
13.1556 | -105.4547
13.1592 | -105.4551
13.1628 | -105.4498
13.1664 | -105.4456
13.1701 | -105.4397
13.1737 | -105.4316
13.1773 | -105.4221
13.181 | -105.4071
13.1846 | -105.3953
13.1882 | -105.3858
13.1918 | -105.3763
13.1955 | -105.3709
13.1991 | -105.3673
13.2027 | -105.3646
13.2063 | -105.3589
13.21 | -105.3498
13.2136 | -105.3462
13.2172 | -105.3447
13.2209 | -105.3426
13.2245 | -105.3431
13.2281 | -105.342
13.2317 | -105.3377
13.2354 | -105.3319
13.239 | -105.326
13.2426 | -105.321
13.2463 | -105.3139
13.2499 | -105.3054
13.2535 | -105.2971
13.2571 | -105.2903
13.2608 | -105.2829
13.2644 | -105.2733
13.268 | -105.2627
13.2717 | -105.2514
13.2753 | -105.2385
13.2789 | -105.2244
13.2825 | -105.2086
13.2862 | -105.1941
13.2898 | -105.1785
13.2934 | -105.158
13.2971 | -105.1361
13.3007 | -105.1145
13.3043 | -105.0981
13.3079 | -105.0807
13.3116 | -105.061
13.3152 | -105.0424
13.3188 | -105.0228
13.3224 | -105.0018
13.3261 | -104.9824
13.3297 | -104.9647
13.3333 | -104.9472
13.337 | -104.9285
13.3406 | -104.9123
13.3442 | -104.9004
13.3478 | -104.8861
13.3515 | -104.8682
13.3551 | -104.8584
13.3587 | -104.8479
13.3624 | -104.8403
13.366 | -104.8319
13.3696 | -104.8228
13.3732 | -104.8148
13.3769 | -104.807
13.3805 | -104.8003
13.3841 | -104.7933
13.3878 | -104.7846
13.3914 | -104.778
13.395 | -104.7707
13.3986 | -104.767
13.4023 | -104.7667
13.4059 | -104.761
13.4095 | -104.7557
13.4132 | -104.7516
13.4168 | -104.7482
13.4204 | -104.7504
13.424 | -104.7546
13.4277 | -104.7592
13.4313 | -104.7639
13.4349 | -104.7678
13.4385 | -104.777
13.4422 | -104.787
13.4458 | -104.7951
13.4494 | -104.8058
13.4531 | -104.8194
13.4567 | -104.8364
13.4603 | -104.8549
13.4639 | -104.8747
13.4676 | -104.8997
13.4712 | -104.9267
13.4748 | -104.9532
13.4785 | -104.9788
13.4821 | -105.0067
13.4857 | -105.0397
13.4893 | -105.0726
13.493 | -105.1048
13.4966 | -105.1337
13.5002 | -105.1633
13.5039 | -105.1944
13.5075 | -105.2273
13.5111 | -105.2615
13.5147 | -105.295
13.5184 | -105.328
13.522 | -105.3612
13.5256 | -105.3931
13.5293 | -105.4269
13.5329 | -105.4567
13.5365 | -105.484
13.5401 | -105.5116
13.5438 | -105.5383
13.5474 | -105.5584
13.551 | -105.5786
13.5546 | -105.5935
13.5583 | -105.6086
13.5619 | -105.6258
13.5655 | -105.6417
13.5692 | -105.6565
13.5728 | -105.6715
13.5764 | -105.6853
13.58 | -105.6983
13.5837 | -105.7079
13.5873 | -105.7204
13.5909 | -105.7311
13.5946 | -105.7413
13.5982 | -105.7525
13.6018 | -105.761
13.6054 | -105.7678
13.6091 | -105.776
13.6127 | -105.7835
13.6163 | -105.7862
13.62 | -105.787
13.6236 | -105.7922
13.6272 | -105.7979
13.6308 | -105.8017
13.6345 | -105.8001
13.6381 | -105.8008
13.6417 | -105.8047
13.6454 | -105.8048
13.649 | -105.8045
13.6526 | -105.7975
13.6562 | -105.794
13.6599 | -105.7914
13.6635 | -105.7856
13.6671 | -105.781
13.6707 | -105.7791
13.6744 | -105.7775
13.678 | -105.7707
13.6816 | -105.7577
13.6853 | -105.746
13.6889 | -105.7364
13.6925 | -105.7271
13.6961 | -105.7162
13.6998 | -105.7067
13.7034 | -105.6957
13.707 | -105.6802
13.7107 | -105.6695
13.7143 | -105.6589
13.7179 | -105.6473
13.7215 | -105.6354
13.7252 | -105.6246
13.7288 | -105.6197
13.7324 | -105.6152
13.7361 | -105.6103
13.7397 | -105.605
13.7433 | -105.6007
13.7469 | -105.5961
13.7506 | -105.5946
13.7542 | -105.5939
13.7578 | -105.5934
13.7615 | -105.5919
13.7651 | -105.59
13.7687 | -105.5864
13.7723 | -105.5854
13.776 | -105.5877
13.7796 | -105.5916
13.7832 | -105.5912
13.7868 | -105.5914
13.7905 | -105.5905
13.7941 | -105.593
13.7977 | -105.5978
13.8014 | -105.602
13.805 | -105.6049
13.8086 | -105.6091
13.8122 | -105.615
13.8159 | -105.6207
13.8195 | -105.6251
13.8231 | -105.6268
13.8268 | -105.6272
13.8304 | -105.6295
13.834 | -105.6315
13.8376 | -105.6301
13.8413 | -105.6273
13.8449 | -105.6223
13.8485 | -105.6177
13.8522 | -105.6163
13.8558 | -105.6129
13.8594 | -105.6068
13.863 | -105.6006
13.8667 | -105.5943
13.8703 | -105.586
13.8739 | -105.5792
13.8776 | -105.5776
13.8812 | -105.5787
13.8848 | -105.5771
13.8884 | -105.5732
13.8921 | -105.5683
13.8957 | -105.564
13.8993 | -105.562
13.9029 | -105.5625
13.9066 | -105.5653
13.9102 | -105.5663
13.9138 | -105.5671
13.9175 | -105.5697
13.9211 | -105.5737
13.9247 | -105.5765
13.9283 | -105.5756
13.932 | -105.5773
13.9356 | -105.5807
13.9392 | -105.586
13.9429 | -105.5902
13.9465 | -105.593
13.9501 | -105.5973
13.9537 | -105.601
13.9574 | -105.5996
13.961 | -105.5909
13.9646 | -105.5794
13.9683 | -105.5719
13.9719 | -105.567
13.9755 | -105.5599
13.9791 | -105.5508
13.9828 | -105.544
13.9864 | -105.5366
13.99 | -105.5284
13.9937 | -105.5192
13.9973 | -105.5101
14.0009 | -105.5022
14.0045 | -105.4947
14.0082 | -105.4873
14.0118 | -105.4812
14.0154 | -105.4776
14.019 | -105.4754
14.0227 | -105.4742
14.0263 | -105.4727
14.0299 | -105.4734
14.0336 | -105.4774
14.0372 | -105.4824
14.0408 | -105.4887
14.0444 | -105.5
14.0481 | -105.5118
14.0517 | -105.5235
14.0553 | -105.5363
14.059 | -105.5525
14.0626 | -105.5702
14.0662 | -105.5886
14.0698 | -105.6134
14.0735 | -105.6361
14.0771 | -105.658
14.0807 | -105.6823
14.0844 | -105.7076
14.088 | -105.7297
14.0916 | -105.7537
14.0952 | -105.7759
14.0989 | -105.7951
14.1025 | -105.8123
14.1061 | -105.8302
14.1098 | -105.8464
14.1134 | -105.8592
14.117 | -105.8672
14.1206 | -105.8716
14.1243 | -105.876
14.1279 | -105.8819
14.1315 | -105.8758
14.1351 | -105.8754
14.1388 | -105.8738
14.1424 | -105.8692
14.146 | -105.8679
14.1497 | -105.868
14.1533 | -105.8667
14.1569 | -105.8635
14.1605 | -105.8641
14.1642 | -105.8716
14.1678 | -105.8776
14.1714 | -105.8785
14.1751 | -105.874
14.1787 | -105.874
14.1823 | -105.8784
14.1859 | -105.8832
14.1896 | -105.8901
14.1932 | -105.8986
14.1968 | -105.9081
14.2005 | -105.9214
14.2041 | -105.9354
14.2077 | -105.9512
14.2113 | -105.9696
14.215 | -105.9856
14.2186 | -106.003
14.2222 | -106.0221
14.2258 | -106.042
14.2295 | -106.0636
14.2331 | -106.0827
14.2367 | -106.0991
14.2404 | -106.1158
14.244 | -106.1336
14.2476 | -106.1482
14.2512 | -106.1596
14.2549 | -106.1689
14.2585 | -106.1755
14.2621 | -106.1794
14.2658 | -106.1788
14.2694 | -106.1777
14.273 | -106.1724
14.2766 | -106.1656
14.2803 | -106.1533
14.2839 | -106.1376
14.2875 | -106.1217
14.2912 | -106.1041
14.2948 | -106.0831
14.2984 | -106.0619
14.302 | -106.0388
14.3057 | -106.0143
14.3093 | -105.9863
14.3129 | -105.9563
14.3166 | -105.9271
14.3202 | -105.9016
14.3238 | -105.8768
14.3274 | -105.8489
14.3311 | -105.82
14.3347 | -105.7919
14.3383 | -105.7673
14.342 | -105.7392
14.3456 | -105.7145
14.3492 | -105.6929
14.3528 | -105.6698
14.3565 | -105.6488
14.3601 | -105.6291
14.3637 | -105.6143
14.3673 | -105.6011
14.371 | -105.5884
14.3746 | -105.5785
14.3782 | -105.5723
14.3819 | -105.5682
14.3855 | -105.5634
14.3891 | -105.5604
14.3927 | -105.5618
14.3964 | -105.565
14.4 | -105.5698
14.4036 | -105.5773
14.4073 | -105.5843
14.4109 | -105.5955
14.4145 | -105.6105
14.4181 | -105.627
14.4218 | -105.6461
14.4254 | -105.6664
14.429 | -105.6873
14.4327 | -105.7086
14.4363 | -105.7318
14.4399 | -105.7542
14.4435 | -105.7794
14.4472 | -105.8049
14.4508 | -105.8314
14.4544 | -105.8587
14.4581 | -105.886
14.4617 | -105.9144
14.4653 | -105.9428
14.4689 | -105.9713
14.4726 | -106.002
14.4762 | -106.0285
14.4798 | -106.0554
14.4834 | -106.0818
14.4871 | -106.1071
14.4907 | -106.1288
14.4943 | -106.15
14.498 | -106.1726
14.5016 | -106.1924
14.5052 | -106.2076
14.5088 | -106.2186
14.5125 | -106.2288
14.5161 | -106.2383
14.5197 | -106.2459
14.5234 | -106.2503
14.527 | -106.2505
14.5306 | -106.2465
14.5342 | -106.2404
14.5379 | -106.2326
14.5415 | -106.2257
14.5451 | -106.2173
14.5488 | -106.2089
14.5524 | -106.1958
14.556 | -106.1806
14.5596 | -106.1659
14.5633 | -106.1525
14.5669 | -106.1399
14.5705 | -106.1251
14.5741 | -106.1133
14.5778 | -106.105
14.5814 | -106.0965
14.585 | -106.0851
14.5887 | -106.0763
14.5923 | -106.0689
14.5959 | -106.0616
14.5995 | -106.0533
14.6032 | -106.0463
14.6068 | -106.0451
14.6104 | -106.0453
14.6141 | -106.0487
14.6177 | -106.0525
14.6213 | -106.0564
14.6249 | -106.063
14.6286 | -106.0708
14.6322 | -106.0783
14.6358 | -106.0864
14.6395 | -106.0965
14.6431 | -106.1051
14.6467 | -106.1185
14.6503 | -106.1314
14.654 | -106.1393
14.6576 | -106.1487
14.6612 | -106.1571
14.6649 | -106.1657
14.6685 | -106.1738
14.6721 | -106.1803
14.6757 | -106.1866
14.6794 | -106.1928
14.683 | -106.1954
14.6866 | -106.1994
14.6902 | -106.2009
14.6939 | -106.1999
14.6975 | -106.2004
14.7011 | -106.2021
14.7048 | -106.1985
14.7084 | -106.195
14.712 | -106.1917
14.7156 | -106.1896
14.7193 | -106.1894
14.7229 | -106.1885
14.7265 | -106.1862
14.7302 | -106.1858
14.7338 | -106.187
14.7374 | -106.1893
14.741 | -106.1951
14.7447 | -106.2019
14.7483 | -106.2077
14.7519 | -106.2141
14.7556 | -106.2221
14.7592 | -106.2274
14.7628 | -106.2349
14.7664 | -106.2422
14.7701 | -106.2502
14.7737 | -106.257
14.7773 | -106.2637
14.781 | -106.27
14.7846 | -106.2726
14.7882 | -106.2737
14.7918 | -106.2758
14.7955 | -106.2765
14.7991 | -106.2749
14.8027 | -106.2713
14.8063 | -106.2651
14.81 | -106.2568
14.8136 | -106.2469
14.8172 | -106.2342
14.8209 | -106.218
14.8245 | -106.2007
14.8281 | -106.1832
14.8317 | -106.1643
14.8354 | -106.1444
14.839 | -106.1279
14.8426 | -106.1122
14.8463 | -106.0961
14.8499 | -106.0788
14.8535 | -106.063
14.8571 | -106.0445
14.8608 | -106.0274
14.8644 | -106.0132
14.868 | -106.0027
14.8717 | -105.9913
14.8753 | -105.9776
14.8789 | -105.9691
14.8825 | -105.9603
14.8862 | -105.9509
14.8898 | -105.9452
14.8934 | -105.9398
14.8971 | -105.9377
14.9007 | -105.9343
14.9043 | -105.9317
14.9079 | -105.9317
14.9116 | -105.9285
14.9152 | -105.9243
14.9188 | -105.9219
14.9224 | -105.9172
14.9261 | -105.9119
14.9297 | -105.9069
14.9333 | -105.8996
14.937 | -105.8921
14.9406 | -105.8863
14.9442 | -105.8807
14.9478 | -105.8734
14.9515 | -105.8624
14.9551 | -105.851
14.9587 | -105.8407
14.9624 | -105.8312
14.966 | -105.8194
14.9696 | -105.8095
14.9732 | -105.8033
14.9769 | -105.7948
14.9805 | -105.7882
14.9841 | -105.7798
14.9878 | -105.7704
14.9914 | -105.7615
14.995 | -105.7543
14.9986 | -105.7493
15.0023 | -105.7468
15.0059 | -105.7457
15.0095 | -105.7466
15.0132 | -105.749
15.0168 | -105.7527
15.0204 | -105.7582
15.024 | -105.7667
15.0277 | -105.7742
15.0313 | -105.7807
15.0349 | -105.7888
15.0385 | -105.7982
15.0422 | -105.8093
15.0458 | -105.8185
15.0494 | -105.8269
15.0531 | -105.8349
15.0567 | -105.8437
15.0603 | -105.8529
15.0639 | -105.8615
15.0676 | -105.8702
15.0712 | -105.878
15.0748 | -105.8847
15.0785 | -105.889
15.0821 | -105.895
15.0857 | -105.9015
15.0893 | -105.9065
15.093 | -105.907
15.0966 | -105.9081
15.1002 | -105.9126
15.1039 | -105.9135
15.1075 | -105.9114
15.1111 | -105.906
15.1147 | -105.9032
15.1184 | -105.9004
15.122 | -105.8968
15.1256 | -105.8931
15.1293 | -105.889
15.1329 | -105.8822
15.1365 | -105.8771
15.1401 | -105.8698
15.1438 | -105.8595
15.1474 | -105.8511
15.151 | -105.8439
15.1546 | -105.8354
15.1583 | -105.8265
15.1619 | -105.8199
15.1655 | -105.8151
15.1692 | -105.8101
15.1728 | -105.8064
15.1764 | -105.8041
15.18 | -105.8026
15.1837 | -105.8018
15.1873 | -105.7996
15.1909 | -105.7996
15.1946 | -105.7996
15.1982 | -105.7972
15.2018 | -105.7982
15.2054 | -105.8007
15.2091 | -105.8049
15.2127 | -105.8089
15.2163 | -105.811
15.22 | -105.8093
15.2236 | -105.809
15.2272 | -105.8097
15.2308 | -105.8088
15.2345 | -105.8059
15.2381 | -105.8014
15.2417 | -105.7958
15.2454 | -105.7918
15.249 | -105.788
15.2526 | -105.7857
15.2562 | -105.7834
15.2599 | -105.78
15.2635 | -105.7773
15.2671 | -105.7741
15.2707 | -105.771
15.2744 | -105.767
15.278 | -105.7619
15.2816 | -105.7578
15.2853 | -105.7556
15.2889 | -105.7533
15.2925 | -105.7505
15.2961 | -105.7453
15.2998 | -105.7392
15.3034 | -105.7376
15.307 | -105.7397
15.3107 | -105.7391
15.3143 | -105.7398
15.3179 | -105.7394
15.3215 | -105.7378
15.3252 | -105.7392
15.3288 | -105.743
15.3324 | -105.7452
15.3361 | -105.7481
15.3397 | -105.7539
15.3433 | -105.7615
15.3469 | -105.7664
15.3506 | -105.7724
15.3542 | -105.7779
15.3578 | -105.7851
15.3615 | -105.7924
15.3651 | -105.7975
15.3687 | -105.801
15.3723 | -105.8047
15.376 | -105.8079
15.3796 | -105.8091
15.3832 | -105.81
15.3868 | -105.8121
15.3905 | -105.8106
15.3941 | -105.8112
15.3977 | -105.8105
15.4014 | -105.8042
15.405 | -105.797
15.4086 | -105.7902
15.4122 | -105.7813
15.4159 | -105.773
15.4195 | -105.7672
15.4231 | -105.7608
15.4268 | -105.7535
15.4304 | -105.7413
15.434 | -105.7301
15.4376 | -105.7213
15.4413 | -105.7149
15.4449 | -105.7069
15.4485 | -105.6986
15.4522 | -105.6948
15.4558 | -105.6912
15.4594 | -105.688
15.463 | -105.6862
15.4667 | -105.6845
15.4703 | -105.6836
15.4739 | -105.6808
15.4776 | -105.6803
15.4812 | -105.6843
15.4848 | -105.6903
15.4884 | -105.6986
15.4921 | -105.7073
15.4957 | -105.7179
15.4993 | -105.73
15.5029 | -105.741
15.5066 | -105.7504
15.5102 | -105.762
15.5138 | -105.7761
15.5175 | -105.7913
15.5211 | -105.8074
15.5247 | -105.8194
15.5283 | -105.8281
15.532 | -105.837
15.5356 | -105.847
15.5392 | -105.8539
15.5429 | -105.8636
15.5465 | -105.8713
15.5501 | -105.8757
15.5537 | -105.8784
15.5574 | -105.8799
15.561 | -105.8746
15.5646 | -105.8721
15.5683 | -105.8698
15.5719 | -105.8645
15.5755 | -105.8576
15.5791 | -105.8515
15.5828 | -105.844
15.5864 | -105.8347
15.59 | -105.8242
15.5937 | -105.813
15.5973 | -105.8017
15.6009 | -105.7919
15.6045 | -105.7835
15.6082 | -105.7743
15.6118 | -105.7675
15.6154 | -105.7559
15.619 | -105.7484
15.6227 | -105.7448
15.6263 | -105.7425
15.6299 | -105.7393
15.6336 | -105.7379
15.6372 | -105.7395
15.6408 | -105.7425
15.6444 | -105.7466
15.6481 | -105.7504
15.6517 | -105.7555
15.6553 | -105.7619
15.659 | -105.7701
15.6626 | -105.778
15.6662 | -105.788
15.6698 | -105.801
15.6735 | -105.8112
15.6771 | -105.8219
15.6807 | -105.8319
15.6844 | -105.8429
15.688 | -105.8525
15.6916 | -105.8625
15.6952 | -105.8759
15.6989 | -105.8844
15.7025 | -105.8879
15.7061 | -105.8938
15.7098 | -105.9015
15.7134 | -105.9072
15.717 | -105.9099
15.7206 | -105.9135
15.7243 | -105.9183
15.7279 | -105.9214
15.7315 | -105.9232
15.7351 | -105.9242
15.7388 | -105.924
15.7424 | -105.9236
15.746 | -105.9205
15.7497 | -105.9181
15.7533 | -105.9198
15.7569 | -105.9216
15.7605 | -105.9193
15.7642 | -105.9179
15.7678 | -105.9118
15.7714 | -105.9082
15.7751 | -105.9058
15.7787 | -105.9025
15.7823 | -105.8987
15.7859 | -105.8984
15.7896 | -105.8927
15.7932 | -105.8892
15.7968 | -105.8881
15.8005 | -105.8845
15.8041 | -105.88
15.8077 | -105.8788
15.8113 | -105.8777
15.815 | -105.8783
15.8186 | -105.8788
15.8222 | -105.8799
15.8259 | -105.8811
15.8295 | -105.8817
15.8331 | -105.8841
15.8367 | -105.8882
15.8404 | -105.8976
15.844 | -105.9036
15.8476 | -105.9082
15.8512 | -105.9114
15.8549 | -105.9142
15.8585 | -105.9192
15.8621 | -105.9227
15.8658 | -105.9257
15.8694 | -105.9288
15.873 | -105.9308
15.8766 | -105.9352
15.8803 | -105.939
15.8839 | -105.9398
15.8875 | -105.9413
15.8912 | -105.9398
15.8948 | -105.936
15.8984 | -105.9342
15.902 | -105.9321
15.9057 | -105.928
15.9093 | -105.922
15.9129 | -105.9145
15.9166 | -105.9026
15.9202 | -105.8909
15.9238 | -105.882
15.9274 | -105.8728
15.9311 | -105.8596
15.9347 | -105.8459
15.9383 | -105.8294
15.942 | -105.8133
15.9456 | -105.7977
15.9492 | -105.78
15.9528 | -105.7605
15.9565 | -105.7456
15.9601 | -105.7329
15.9637 | -105.7182
15.9673 | -105.7035
15.971 | -105.6915
15.9746 | -105.6805
15.9782 | -105.6686
15.9819 | -105.6565
15.9855 | -105.646
15.9891 | -105.6389
15.9927 | -105.635
15.9964 | -105.6348
16 | -105.6341
16.0036 | -105.6341
16.0073 | -105.6358
16.0109 | -105.6388
16.0145 | -105.6431
16.0181 | -105.6514
16.0218 | -105.662
16.0254 | -105.6733
16.029 | -105.685
16.0327 | -105.6982
16.0363 | -105.7123
16.0399 | -105.7287
16.0435 | -105.7443
16.0472 | -105.7611
16.0508 | -105.7779
16.0544 | -105.795
16.0581 | -105.812
16.0617 | -105.8287
16.0653 | -105.8474
16.0689 | -105.8652
16.0726 | -105.8819
16.0762 | -105.8977
16.0798 | -105.9122
16.0834 | -105.9297
16.0871 | -105.9443
16.0907 | -105.9558
16.0943 | -105.9659
16.098 | -105.9768
16.1016 | -105.988
16.1052 | -105.9959
16.1088 | -106.0022
16.1125 | -106.0075
16.1161 | -106.0089
16.1197 | -106.0122
16.1234 | -106.0168
16.127 | -106.0182
16.1306 | -106.018
16.1342 | -106.0175
16.1379 | -106.0157
16.1415 | -106.0181
16.1451 | -106.021
16.1488 | -106.0196
16.1524 | -106.0167
16.156 | -106.013
16.1596 | -106.0116
16.1633 | -106.0097
16.1669 | -106.011
16.1705 | -106.0133
16.1742 | -106.0165
16.1778 | -106.0198
16.1814 | -106.0214
16.185 | -106.0214
16.1887 | -106.0202
16.1923 | -106.0202
16.1959 | -106.0209
16.1995 | -106.0204
16.2032 | -106.0206
16.2068 | -106.0228
16.2104 | -106.0234
16.2141 | -106.025
16.2177 | -106.0238
16.2213 | -106.0223
16.2249 | -106.0228
16.2286 | -106.02
16.2322 | -106.0178
16.2358 | -106.0192
16.2395 | -106.0207
16.2431 | -106.023
16.2467 | -106.0245
16.2503 | -106.0252
16.254 | -106.0226
16.2576 | -106.0227
16.2612 | -106.0252
16.2649 | -106.0249
16.2685 | -106.0218
16.2721 | -106.0205
16.2757 | -106.0201
16.2794 | -106.0184
16.283 | -106.0178
16.2866 | -106.0164
16.2902 | -106.0158
16.2939 | -106.016
16.2975 | -106.0129
16.3011 | -106.0098
16.3048 | -106.0098
16.3084 | -106.0102
16.312 | -106.0122
16.3156 | -106.0117
16.3193 | -106.01
16.3229 | -106.0128
16.3265 | -106.016
16.3302 | -106.016
16.3338 | -106.0181
16.3374 | -106.0217
16.341 | -106.025
16.3447 | -106.0287
16.3483 | -106.0343
16.3519 | -106.036
16.3556 | -106.0383
16.3592 | -106.0413
16.3628 | -106.0468
16.3664 | -106.0538
16.3701 | -106.0613
16.3737 | -106.0658
16.3773 | -106.0729
16.381 | -106.0805
16.3846 | -106.0881
16.3882 | -106.096
16.3918 | -106.1047
16.3955 | -106.1153
16.3991 | -106.1244
16.4027 | -106.1345
16.4063 | -106.1439
16.41 | -106.1537
16.4136 | -106.1673
16.4172 | -106.18
16.4209 | -106.1912
16.4245 | -106.2035
16.4281 | -106.2167
16.4317 | -106.2291
16.4354 | -106.2398
16.439 | -106.2491
16.4426 | -106.26
16.4463 | -106.2704
16.4499 | -106.2796
16.4535 | -106.2879
16.4571 | -106.295
16.4608 | -106.3028
16.4644 | -106.3117
16.468 | -106.3186
16.4717 | -106.3224
16.4753 | -106.3244
16.4789 | -106.3251
16.4825 | -106.3269
16.4862 | -106.3291
16.4898 | -106.3304
16.4934 | -106.3296
16.4971 | -106.329
16.5007 | -106.327
16.5043 | -106.3259
16.5079 | -106.3251
16.5116 | -106.3222
16.5152 | -106.3176
16.5188 | -106.3116
16.5224 | -106.3074
16.5261 | -106.3046
16.5297 | -106.2986
16.5333 | -106.2951
16.537 | -106.2926
16.5406 | -106.2895
16.5442 | -106.2875
16.5478 | -106.2843
16.5515 | -106.2823
16.5551 | -106.2813
16.5587 | -106.2782
16.5624 | -106.276
16.566 | -106.2738
16.5696 | -106.2759
16.5732 | -106.2793
16.5769 | -106.2825
16.5805 | -106.2868
16.5841 | -106.292
16.5878 | -106.3002
16.5914 | -106.3077
16.595 | -106.3168
16.5986 | -106.3272
16.6023 | -106.3367
16.6059 | -106.3461
16.6095 | -106.358
16.6132 | -106.3708
16.6168 | -106.3844
16.6204 | -106.3986
16.624 | -106.4146
16.6277 | -106.4303
16.6313 | -106.4464
16.6349 | -106.4626
16.6385 | -106.4769
16.6422 | -106.4926
16.6458 | -106.5095
16.6494 | -106.5246
16.6531 | -106.5408
16.6567 | -106.5556
16.6603 | -106.5694
16.6639 | -106.5832
16.6676 | -106.5966
16.6712 | -106.6089
16.6748 | -106.6224
16.6785 | -106.637
16.6821 | -106.6491
16.6857 | -106.6596
16.6893 | -106.6704
16.693 | -106.68
16.6966 | -106.6875
16.7002 | -106.6955
16.7039 | -106.7035
16.7075 | -106.7094
16.7111 | -106.7161
16.7147 | -106.7218
16.7184 | -106.7244
16.722 | -106.7264
16.7256 | -106.73
16.7293 | -106.7332
16.7329 | -106.7335
16.7365 | -106.7355
16.7401 | -106.7368
16.7438 | -106.733
16.7474 | -106.7283
16.751 | -106.727
16.7546 | -106.7274
16.7583 | -106.7243
16.7619 | -106.7189
16.7655 | -106.714
16.7692 | -106.7119
16.7728 | -106.7087
16.7764 | -106.7045
16.78 | -106.6989
16.7837 | -106.6947
16.7873 | -106.6913
16.7909 | -106.6879
16.7946 | -106.6844
16.7982 | -106.6832
16.8018 | -106.6835
16.8054 | -106.6837
16.8091 | -106.6806
16.8127 | -106.6767
16.8163 | -106.6734
16.82 | -106.6721
16.8236 | -106.6706
16.8272 | -106.6688
16.8308 | -106.6651
16.8345 | -106.6621
16.8381 | -106.6574
16.8417 | -106.6556
16.8454 | -106.6545
16.849 | -106.6506
16.8526 | -106.6442
16.8562 | -106.639
16.8599 | -106.6342
16.8635 | -106.627
16.8671 | -106.6178
16.8707 | -106.6072
16.8744 | -106.5965
16.878 | -106.5882
16.8816 | -106.5805
16.8853 | -106.5733
16.8889 | -106.566
16.8925 | -106.5596
16.8961 | -106.5544
16.8998 | -106.5476
16.9034 | -106.5391
16.907 | -106.5297
16.9107 | -106.5235
16.9143 | -106.5157
16.9179 | -106.5053
16.9215 | -106.4974
16.9252 | -106.4933
16.9288 | -106.4865
16.9324 | -106.4772
16.9361 | -106.4703
16.9397 | -106.4638
16.9433 | -106.4563
16.9469 | -106.4503
16.9506 | -106.4461
16.9542 | -106.4397
16.9578 | -106.4344
16.9615 | -106.4289
16.9651 | -106.4235
16.9687 | -106.4176
16.9723 | -106.4119
16.976 | -106.4076
16.9796 | -106.4033
16.9832 | -106.4005
16.9868 | -106.3969
16.9905 | -106.393
16.9941 | -106.3872
16.9977 | -106.3842
17.0014 | -106.3837
17.005 | -106.3815
17.0086 | -106.3785
17.0122 | -106.377
17.0159 | -106.3767
17.0195 | -106.3782
17.0231 | -106.3797
17.0268 | -106.3817
17.0304 | -106.3834
17.034 | -106.384
17.0376 | -106.387
17.0413 | -106.3876
17.0449 | -106.3893
17.0485 | -106.3953
17.0522 | -106.3997
17.0558 | -106.4035
17.0594 | -106.4076
17.063 | -106.4114
17.0667 | -106.415
17.0703 | -106.4172
17.0739 | -106.4203
17.0776 | -106.423
17.0812 | -106.4258
17.0848 | -106.4251
17.0884 | -106.4228
17.0921 | -106.4232
17.0957 | -106.4223
17.0993 | -106.418
17.1029 | -106.4146
17.1066 | -106.4131
17.1102 | -106.4103
17.1138 | -106.4063
17.1175 | -106.4025
17.1211 | -106.3966
17.1247 | -106.3884
17.1283 | -106.3793
17.132 | -106.3709
17.1356 | -106.3606
17.1392 | -106.3511
17.1429 | -106.3426
17.1465 | -106.3335
17.1501 | -106.3254
17.1537 | -106.314
17.1574 | -106.3059
17.161 | -106.2982
17.1646 | -106.2875
17.1683 | -106.2765
17.1719 | -106.2658
17.1755 | -106.258
17.1791 | -106.2491
17.1828 | -106.2419
17.1864 | -106.2354
17.19 | -106.229
17.1937 | -106.2219
17.1973 | -106.2174
17.2009 | -106.2132
17.2045 | -106.2083
17.2082 | -106.2038
17.2118 | -106.2008
17.2154 | -106.2008
17.219 | -106.1998
17.2227 | -106.1983
17.2263 | -106.196
17.2299 | -106.1938
17.2336 | -106.194
17.2372 | -106.1959
17.2408 | -106.1964
17.2444 | -106.1965
17.2481 | -106.1967
17.2517 | -106.197
17.2553 | -106.1951
17.259 | -106.1922
17.2626 | -106.1889
17.2662 | -106.189
17.2698 | -106.1838
17.2735 | -106.1799
17.2771 | -106.1784
17.2807 | -106.1759
17.2844 | -106.1708
17.288 | -106.1687
17.2916 | -106.1649
17.2952 | -106.1595
17.2989 | -106.1533
17.3025 | -106.1466
17.3061 | -106.1401
17.3098 | -106.1341
17.3134 | -106.13
17.317 | -106.1282
17.3206 | -106.1245
17.3243 | -106.117
17.3279 | -106.1094
17.3315 | -106.105
17.3351 | -106.103
17.3388 | -106.0977
17.3424 | -106.0937
17.346 | -106.092
17.3497 | -106.0897
17.3533 | -106.0876
17.3569 | -106.0869
17.3605 | -106.0869
17.3642 | -106.0824
17.3678 | -106.0779
17.3714 | -106.0759
17.3751 | -106.0747
17.3787 | -106.0723
17.3823 | -106.0722
17.3859 | -106.0731
17.3896 | -106.0736
17.3932 | -106.0766
17.3968 | -106.0812
17.4005 | -106.0848
17.4041 | -106.0856
17.4077 | -106.0877
17.4113 | -106.0919
17.415 | -106.0966
17.4186 | -106.1031
17.4222 | -106.1106
17.4259 | -106.1167
17.4295 | -106.1208
17.4331 | -106.1256
17.4367 | -106.1305
17.4404 | -106.1362
17.444 | -106.1434
17.4476 | -106.1483
17.4512 | -106.152
17.4549 | -106.156
17.4585 | -106.1613
17.4621 | -106.1646
17.4658 | -106.1666
17.4694 | -106.1669
17.473 | -106.1676
17.4766 | -106.1703
17.4803 | -106.1706
17.4839 | -106.168
17.4875 | -106.1647
17.4912 | -106.1647
17.4948 | -106.1634
17.4984 | -106.1587
17.502 | -106.1566
17.5057 | -106.1557
17.5093 | -106.1522
17.5129 | -106.1443
17.5166 | -106.1356
17.5202 | -106.1286
17.5238 | -106.1224
17.5274 | -106.1146
17.5311 | -106.107
17.5347 | -106.0982
17.5383 | -106.0873
17.5419 | -106.0767
17.5456 | -106.0662
17.5492 | -106.0538
17.5528 | -106.0461
17.5565 | -106.0395
17.5601 | -106.0292
17.5637 | -106.0159
17.5673 | -106.0016
17.571 | -105.99
17.5746 | -105.9826
17.5782 | -105.9748
17.5819 | -105.9646
17.5855 | -105.955
17.5891 | -105.9476
17.5927 | -105.9397
17.5964 | -105.9293
17.6 | -105.9189
17.6036 | -105.9085
17.6073 | -105.8983
17.6109 | -105.8881
17.6145 | -105.8828
17.6181 | -105.8772
17.6218 | -105.868
17.6254 | -105.8592
17.629 | -105.85
17.6327 | -105.8398
17.6363 | -105.8302
17.6399 | -105.8215
17.6435 | -105.8108
17.6472 | -105.8014
17.6508 | -105.7911
17.6544 | -105.7815
17.658 | -105.7712
17.6617 | -105.7629
17.6653 | -105.7545
17.6689 | -105.7467
17.6726 | -105.7382
17.6762 | -105.7288
17.6798 | -105.7196
17.6834 | -105.7115
17.6871 | -105.7027
17.6907 | -105.6933
17.6943 | -105.6832
17.698 | -105.6755
17.7016 | -105.67
17.7052 | -105.6647
17.7088 | -105.6571
17.7125 | -105.6519
17.7161 | -105.646
17.7197 | -105.6429
17.7234 | -105.6393
17.727 | -105.6329
17.7306 | -105.6262
17.7342 | -105.6239
17.7379 | -105.6219
17.7415 | -105.6184
17.7451 | -105.617
17.7488 | -105.6147
17.7524 | -105.6115
17.756 | -105.6094
17.7596 | -105.6072
17.7633 | -105.6074
17.7669 | -105.608
17.7705 | -105.6098
17.7741 | -105.6107
17.7778 | -105.6119
17.7814 | -105.6116
17.785 | -105.6149
17.7887 | -105.6177
17.7923 | -105.6175
17.7959 | -105.6165
17.7995 | -105.6182
17.8032 | -105.6223
17.8068 | -105.62
17.8104 | -105.6156
17.8141 | -105.6152
17.8177 | -105.6167
17.8213 | -105.6189
17.8249 | -105.6164
17.8286 | -105.613
17.8322 | -105.6109
17.8358 | -105.61
17.8395 | -105.6073
17.8431 | -105.6056
17.8467 | -105.6049
17.8503 | -105.6032
17.854 | -105.5998
17.8576 | -105.5949
17.8612 | -105.5921
17.8649 | -105.5912
17.8685 | -105.587
17.8721 | -105.5807
17.8757 | -105.5746
17.8794 | -105.5682
17.883 | -105.5632
17.8866 | -105.5566
17.8902 | -105.5524
17.8939 | -105.5471
17.8975 | -105.5416
17.9011 | -105.5326
17.9048 | -105.525
17.9084 | -105.5162
17.912 | -105.5116
17.9156 | -105.5056
17.9193 | -105.4985
17.9229 | -105.4896
17.9265 | -105.4817
17.9302 | -105.4739
17.9338 | -105.4664
17.9374 | -105.4601
17.941 | -105.4538
17.9447 | -105.4475
17.9483 | -105.4415
17.9519 | -105.4349
17.9556 | -105.4254
17.9592 | -105.413
17.9628 | -105.4032
17.9664 | -105.394
17.9701 | -105.3862
17.9737 | -105.3768
17.9773 | -105.3668
17.981 | -105.3564
17.9846 | -105.3459
17.9882 | -105.3338
17.9918 | -105.3217
17.9955 | -105.3066
17.9991 | -105.295
18.0027 | -105.2871
18.0063 | -105.2764
18.01 | -105.2646
18.0136 | -105.2538
18.0172 | -105.244
18.0209 | -105.2363
18.0245 | -105.2275
18.0281 | -105.217
18.0317 | -105.2068
18.0354 | -105.1982
18.039 | -105.1892
18.0426 | -105.1834
18.0463 | -105.18
18.0499 | -105.1742
18.0535 | -105.1649
18.0571 | -105.1586
18.0608 | -105.1538
18.0644 | -105.1516
18.068 | -105.149
18.0717 | -105.1424
18.0753 | -105.1355
18.0789 | -105.1309
18.0825 | -105.1285
18.0862 | -105.1262
18.0898 | -105.1219
18.0934 | -105.1186
18.0971 | -105.116
18.1007 | -105.1129
18.1043 | -105.1094
18.1079 | -105.107
18.1116 | -105.1053
18.1152 | -105.1062
18.1188 | -105.106
18.1224 | -105.104
18.1261 | -105.1021
18.1297 | -105.1048
18.1333 | -105.1085
18.137 | -105.1099
18.1406 | -105.1083
18.1442 | -105.1077
18.1478 | -105.1074
18.1515 | -105.1095
18.1551 | -105.1138
18.1587 | -105.1185
18.1624 | -105.1212
18.166 | -105.1242
18.1696 | -105.1245
18.1732 | -105.125
18.1769 | -105.1293
18.1805 | -105.1356
18.1841 | -105.1409
18.1878 | -105.1478
18.1914 | -105.1525
18.195 | -105.1565
18.1986 | -105.1622
18.2023 | -105.1674
18.2059 | -105.1703
18.2095 | -105.1748
18.2132 | -105.1805
18.2168 | -105.1872
18.2204 | -105.1926
18.224 | -105.1933
18.2277 | -105.1947
18.2313 | -105.1954
18.2349 | -105.1961
18.2385 | -105.1971
18.2422 | -105.1964
18.2458 | -105.1949
18.2494 | -105.1942
18.2531 | -105.1938
18.2567 | -105.1866
18.2603 | -105.1781
18.2639 | -105.1733
18.2676 | -105.1691
18.2712 | -105.1643
18.2748 | -105.1566
18.2785 | -105.1479
18.2821 | -105.1404
18.2857 | -105.1365
18.2893 | -105.1323
18.293 | -105.1286
18.2966 | -105.1257
18.3002 | -105.1215
18.3039 | -105.1174
18.3075 | -105.1153
18.3111 | -105.1138
18.3147 | -105.1104
18.3184 | -105.1076
18.322 | -105.1031
18.3256 | -105.1015
18.3293 | -105.1026
18.3329 | -105.1064
18.3365 | -105.1094
18.3401 | -105.1084
18.3438 | -105.1095
18.3474 | -105.111
18.351 | -105.1127
18.3546 | -105.1129
18.3583 | -105.1119
18.3619 | -105.1126
18.3655 | -105.1159
18.3692 | -105.12
18.3728 | -105.1219
18.3764 | -105.1219
18.38 | -105.1222
18.3837 | -105.1202
18.3873 | -105.1208
18.3909 | -105.1202
18.3946 | -105.118
18.3982 | -105.116
18.4018 | -105.1165
18.4054 | -105.1137
18.4091 | -105.1043
18.4127 | -105.0935
18.4163 | -105.0842
18.42 | -105.0762
18.4236 | -105.071
18.4272 | -105.0645
18.4308 | -105.0536
18.4345 | -105.042
18.4381 | -105.0284
18.4417 | -105.0134
18.4454 | -104.9998
18.449 | -104.9854
18.4526 | -104.9723
18.4562 | -104.9575
18.4599 | -104.9398
18.4635 | -104.9252
18.4671 | -104.9111
18.4707 | -104.8974
18.4744 | -104.882
18.478 | -104.8665
18.4816 | -104.8509
18.4853 | -104.8402
18.4889 | -104.828
18.4925 | -104.8167
18.4961 | -104.8039
18.4998 | -104.7916
18.5034 | -104.781
18.507 | -104.7716
18.5107 | -104.7624
18.5143 | -104.7557
18.5179 | -104.7531
18.5215 | -104.7491
18.5252 | -104.7432
18.5288 | -104.7399
18.5324 | -104.737
18.5361 | -104.7337
18.5397 | -104.7301
18.5433 | -104.7255
18.5469 | -104.721
18.5506 | -104.7185
18.5542 | -104.7167
18.5578 | -104.7134
18.5615 | -104.7094
18.5651 | -104.7054
18.5687 | -104.7015
18.5723 | -104.7
18.576 | -104.696
18.5796 | -104.6921
18.5832 | -104.6871
18.5868 | -104.6818
18.5905 | -104.6742
18.5941 | -104.6653
18.5977 | -104.6545
18.6014 | -104.6422
18.605 | -104.6293
18.6086 | -104.6151
18.6122 | -104.6004
18.6159 | -104.5865
18.6195 | -104.5735
18.6231 | -104.5596
18.6268 | -104.541
18.6304 | -104.5207
18.634 | -104.4979
18.6376 | -104.4756
18.6413 | -104.4529
18.6449 | -104.4307
18.6485 | -104.4105
18.6522 | -104.3871
18.6558 | -104.3636
18.6594 | -104.3381
18.663 | -104.3135
18.6667 | -104.2889
18.6703 | -104.2629
18.6739 | -104.24
18.6776 | -104.2164
18.6812 | -104.1929
18.6848 | -104.1691
18.6884 | -104.1434
18.6921 | -104.1185
18.6957 | -104.0963
18.6993 | -104.0748
18.7029 | -104.0538
18.7066 | -104.0356
18.7102 | -104.0154
18.7138 | -103.9969
18.7175 | -103.979
18.7211 | -103.9611
18.7247 | -103.9467
18.7283 | -103.9293
18.732 | -103.914
18.7356 | -103.8998
18.7392 | -103.8882
18.7429 | -103.8803
18.7465 | -103.8719
18.7501 | -103.8638
18.7537 | -103.856
18.7574 | -103.8467
18.761 | -103.8419
18.7646 | -103.8378
18.7683 | -103.8323
18.7719 | -103.8291
18.7755 | -103.8276
18.7791 | -103.827
18.7828 | -103.826
18.7864 | -103.8263
18.79 | -103.828
18.7937 | -103.8288
18.7973 | -103.8314
18.8009 | -103.8367
18.8045 | -103.8407
18.8082 | -103.8455
18.8118 | -103.8504
18.8154 | -103.854
18.819 | -103.8579
18.8227 | -103.8623
18.8263 | -103.8689
18.8299 | -103.8758
18.8336 | -103.8831
18.8372 | -103.8883
18.8408 | -103.8929
18.8444 | -103.8976
18.8481 | -103.9032
18.8517 | -103.9067
18.8553 | -103.909
18.859 | -103.9087
18.8626 | -103.9095
18.8662 | -103.9131
18.8698 | -103.9186
18.8735 | -103.9232
18.8771 | -103.926
18.8807 | -103.9279
18.8844 | -103.9298
18.888 | -103.9307
18.8916 | -103.9334
18.8952 | -103.9366
18.8989 | -103.9378
18.9025 | -103.9379
18.9061 | -103.9366
18.9098 | -103.9336
18.9134 | -103.9267
18.917 | -103.9206
18.9206 | -103.9166
18.9243 | -103.9109
18.9279 | -103.9061
18.9315 | -103.9
18.9351 | -103.8923
18.9388 | -103.8862
18.9424 | -103.8781
18.946 | -103.8682
18.9497 | -103.8563
18.9533 | -103.8468
18.9569 | -103.8374
18.9605 | -103.8281
18.9642 | -103.8199
18.9678 | -103.8129
18.9714 | -103.803
18.9751 | -103.7933
18.9787 | -103.7869
18.9823 | -103.7804
18.9859 | -103.7747
18.9896 | -103.7706
18.9932 | -103.766
18.9968 | -103.7637
19.0005 | -103.7601
19.0041 | -103.7579
19.0077 | -103.7573
19.0113 | -103.751
19.015 | -103.7476
19.0186 | -103.7464
19.0222 | -103.7463
19.0258 | -103.7497
19.0295 | -103.7514
19.0331 | -103.7518
19.0367 | -103.7547
19.0404 | -103.7565
19.044 | -103.7541
19.0476 | -103.7532
19.0512 | -103.7539
19.0549 | -103.7542
19.0585 | -103.7549
19.0621 | -103.7552
19.0658 | -103.7569
19.0694 | -103.7575
19.073 | -103.759
19.0766 | -103.7601
19.0803 | -103.7588
19.0839 | -103.7587
19.0875 | -103.758
19.0912 | -103.7568
19.0948 | -103.754
19.0984 | -103.7517
19.102 | -103.7503
19.1057 | -103.7493
19.1093 | -103.7482
19.1129 | -103.7473
19.1166 | -103.7475
19.1202 | -103.7465
19.1238 | -103.746
19.1274 | -103.7473
19.1311 | -103.7491
19.1347 | -103.7508
19.1383 | -103.751
19.142 | -103.7503
19.1456 | -103.7521
19.1492 | -103.7548
19.1528 | -103.7563
19.1565 | -103.7597
19.1601 | -103.7638
19.1637 | -103.7686
19.1673 | -103.7771
19.171 | -103.7859
19.1746 | -103.7977
19.1782 | -103.8102
19.1819 | -103.8229
19.1855 | -103.837
19.1891 | -103.8499
19.1927 | -103.8607
19.1964 | -103.8732
19.2 | -103.887
19.2036 | -103.9013
19.2073 | -103.9163
19.2109 | -103.9311
19.2145 | -103.9465
19.2181 | -103.9625
19.2218 | -103.9778
19.2254 | -103.9915
19.229 | -104.0074
19.2327 | -104.0235
19.2363 | -104.0368
19.2399 | -104.0483
19.2435 | -104.0627
19.2472 | -104.075
19.2508 | -104.0864
19.2544 | -104.097
19.2581 | -104.1068
19.2617 | -104.1178
19.2653 | -104.1271
19.2689 | -104.1356
19.2726 | -104.1454
19.2762 | -104.1514
19.2798 | -104.1583
19.2834 | -104.1693
19.2871 | -104.1785
19.2907 | -104.1822
19.2943 | -104.1884
19.298 | -104.1945
19.3016 | -104.197
19.3052 | -104.2005
19.3088 | -104.2043
19.3125 | -104.2061
19.3161 | -104.2122
19.3197 | -104.2154
19.3234 | -104.2164
19.327 | -104.2193
19.3306 | -104.2214
19.3342 | -104.2248
19.3379 | -104.2279
19.3415 | -104.229
19.3451 | -104.231
19.3488 | -104.2349
19.3524 | -104.2401
19.356 | -104.2429
19.3596 | -104.2448
19.3633 | -104.2466
19.3669 | -104.2494
19.3705 | -104.2541
19.3742 | -104.2584
19.3778 | -104.2616
19.3814 | -104.2628
19.385 | -104.2661
19.3887 | -104.2714
19.3923 | -104.2746
19.3959 | -104.2797
19.3995 | -104.2849
19.4032 | -104.2892
19.4068 | -104.2948
19.4104 | -104.3002
19.4141 | -104.3059
19.4177 | -104.3121
19.4213 | -104.3181
19.4249 | -104.3246
19.4286 | -104.332
19.4322 | -104.338
19.4358 | -104.3457
19.4395 | -104.3536
19.4431 | -104.3604
19.4467 | -104.3692
19.4503 | -104.3772
19.454 | -104.387
19.4576 | -104.3981
19.4612 | -104.4098
19.4649 | -104.4204
19.4685 | -104.4329
19.4721 | -104.4442
19.4757 | -104.4547
19.4794 | -104.4661
19.483 | -104.478
19.4866 | -104.4895
19.4902 | -104.4993
19.4939 | -104.5082
19.4975 | -104.5197
19.5011 | -104.5298
19.5048 | -104.5378
19.5084 | -104.5461
19.512 | -104.556
19.5156 | -104.5635
19.5193 | -104.5715
19.5229 | -104.5808
19.5265 | -104.5886
19.5302 | -104.5958
19.5338 | -104.6026
19.5374 | -104.6068
19.541 | -104.6108
19.5447 | -104.6152
19.5483 | -104.6187
19.5519 | -104.6242
19.5556 | -104.6268
19.5592 | -104.6284
19.5628 | -104.6291
19.5664 | -104.6348
19.5701 | -104.6373
19.5737 | -104.6385
19.5773 | -104.6428
19.581 | -104.6478
19.5846 | -104.6503
19.5882 | -104.6518
19.5918 | -104.6518
19.5955 | -104.6547
19.5991 | -104.6595
19.6027 | -104.6613
19.6063 | -104.6608
19.61 | -104.6595
19.6136 | -104.6589
19.6172 | -104.6603
19.6209 | -104.6611
19.6245 | -104.6625
19.6281 | -104.6625
19.6317 | -104.6651
19.6354 | -104.6664
19.639 | -104.6675
19.6426 | -104.6695
19.6463 | -104.6686
19.6499 | -104.6635
19.6535 | -104.6573
19.6571 | -104.6515
19.6608 | -104.6477
19.6644 | -104.6423
19.668 | -104.637
19.6717 | -104.6328
19.6753 | -104.6281
19.6789 | -104.6211
19.6825 | -104.6123
19.6862 | -104.6045
19.6898 | -104.5949
19.6934 | -104.5859
19.6971 | -104.5772
19.7007 | -104.565
19.7043 | -104.5518
19.7079 | -104.5401
19.7116 | -104.528
19.7152 | -104.5135
19.7188 | -104.4974
19.7224 | -104.4835
19.7261 | -104.4701
19.7297 | -104.4566
19.7333 | -104.4396
19.737 | -104.4221
19.7406 | -104.4049
19.7442 | -104.3907
19.7478 | -104.3763
19.7515 | -104.3614
19.7551 | -104.3476
19.7587 | -104.333
19.7624 | -104.3177
19.766 | -104.3054
19.7696 | -104.2952
19.7732 | -104.2859
19.7769 | -104.2747
19.7805 | -104.2663
19.7841 | -104.258
19.7878 | -104.2476
19.7914 | -104.2386
19.795 | -104.2339
19.7986 | -104.2314
19.8023 | -104.2292
19.8059 | -104.2281
19.8095 | -104.2289
19.8132 | -104.2283
19.8168 | -104.2275
19.8204 | -104.2302
19.824 | -104.2349
19.8277 | -104.2391
19.8313 | -104.243
19.8349 | -104.2486
19.8385 | -104.2553
19.8422 | -104.263
19.8458 | -104.2678
19.8494 | -104.2743
19.8531 | -104.2809
19.8567 | -104.2888
19.8603 | -104.2968
19.8639 | -104.3062
19.8676 | -104.3174
19.8712 | -104.3288
19.8748 | -104.337
19.8785 | -104.345
19.8821 | -104.3556
19.8857 | -104.3667
19.8893 | -104.3727
19.893 | -104.3805
19.8966 | -104.3886
19.9002 | -104.3961
19.9039 | -104.4034
19.9075 | -104.4084
19.9111 | -104.4159
19.9147 | -104.4209
19.9184 | -104.4262
19.922 | -104.431
19.9256 | -104.433
19.9293 | -104.4357
19.9329 | -104.439
19.9365 | -104.4415
19.9401 | -104.4423
19.9438 | -104.4435
19.9474 | -104.4433
19.951 | -104.4405
19.9546 | -104.4367
19.9583 | -104.4327
19.9619 | -104.4301
19.9655 | -104.4251
19.9692 | -104.4194
19.9728 | -104.4148
19.9764 | -104.4079
19.98 | -104.3992
19.9837 | -104.3916
19.9873 | -104.3824
19.9909 | -104.3701
19.9946 | -104.359
19.9982 | -104.3481
20.0018 | -104.3371
20.0054 | -104.3246
20.0091 | -104.311
20.0127 | -104.2972
20.0163 | -104.2855
20.02 | -104.2746
20.0236 | -104.2638
20.0272 | -104.252
20.0308 | -104.2407
20.0345 | -104.2314
20.0381 | -104.2211
20.0417 | -104.2132
20.0454 | -104.2063
20.049 | -104.1979
20.0526 | -104.1908
20.0562 | -104.1841
20.0599 | -104.1773
20.0635 | -104.1709
20.0671 | -104.1671
20.0707 | -104.1637
20.0744 | -104.1616
20.078 | -104.1592
20.0816 | -104.1558
20.0853 | -104.1537
20.0889 | -104.153
20.0925 | -104.152
20.0961 | -104.1544
20.0998 | -104.159
20.1034 | -104.1617
20.107 | -104.1637
20.1107 | -104.168
20.1143 | -104.1716
20.1179 | -104.1733
20.1215 | -104.1754
20.1252 | -104.1799
20.1288 | -104.1856
20.1324 | -104.1932
20.1361 | -104.1989
20.1397 | -104.2023
20.1433 | -104.2053
20.1469 | -104.2079
20.1506 | -104.2114
20.1542 | -104.2152
20.1578 | -104.2186
20.1615 | -104.2203
20.1651 | -104.2212
20.1687 | -104.2229
20.1723 | -104.224
20.176 | -104.2276
20.1796 | -104.2306
20.1832 | -104.2303
20.1868 | -104.23
20.1905 | -104.2296
20.1941 | -104.2289
20.1977 | -104.2271
20.2014 | -104.2257
20.205 | -104.2244
20.2086 | -104.2233
20.2122 | -104.2232
20.2159 | -104.2219
20.2195 | -104.22
20.2231 | -104.2186
20.2268 | -104.2152
20.2304 | -104.2083
20.234 | -104.2026
20.2376 | -104.1967
20.2413 | -104.1905
20.2449 | -104.1839
20.2485 | -104.1786
20.2522 | -104.1746
20.2558 | -104.1691
20.2594 | -104.1629
20.263 | -104.1558
20.2667 | -104.1497
20.2703 | -104.1449
20.2739 | -104.1408
20.2776 | -104.1381
20.2812 | -104.133
20.2848 | -104.1272
20.2884 | -104.1213
20.2921 | -104.1144
20.2957 | -104.1084
20.2993 | -104.1048
20.3029 | -104.1012
20.3066 | -104.0966
20.3102 | -104.0924
20.3138 | -104.0877
20.3175 | -104.0823
20.3211 | -104.0768
20.3247 | -104.0691
20.3283 | -104.0619
20.332 | -104.0559
20.3356 | -104.0512
20.3392 | -104.0469
20.3429 | -104.044
20.3465 | -104.0435
20.3501 | -104.0422
20.3537 | -104.0404
20.3574 | -104.0399
20.361 | -104.0378
20.3646 | -104.0355
20.3683 | -104.0344
20.3719 | -104.0344
20.3755 | -104.0338
20.3791 | -104.0334
20.3828 | -104.0329
20.3864 | -104.032
20.39 | -104.0311
20.3937 | -104.0304
20.3973 | -104.0302
20.4009 | -104.0306
20.4045 | -104.034
20.4082 | -104.0367
20.4118 | -104.0347
20.4154 | -104.0336
20.419 | -104.0326
20.4227 | -104.0305
20.4263 | -104.0303
20.4299 | -104.0306
20.4336 | -104.0307
20.4372 | -104.0297
20.4408 | -104.0278
20.4444 | -104.0266
20.4481 | -104.0252
20.4517 | -104.025
20.4553 | -104.0225
20.459 | -104.0192
20.4626 | -104.0171
20.4662 | -104.0128
20.4698 | -104.0096
20.4735 | -104.0073
20.4771 | -104.0038
20.4807 | -103.9987
20.4844 | -103.9953
20.488 | -103.9923
20.4916 | -103.9907
20.4952 | -103.989
20.4989 | -103.9846
20.5025 | -103.9802
20.5061 | -103.9772
20.5098 | -103.9727
20.5134 | -103.9662
20.517 | -103.9598
20.5206 | -103.9566
20.5243 | -103.9535
20.5279 | -103.9484
20.5315 | -103.9421
20.5351 | -103.9378
20.5388 | -103.9323
20.5424 | -103.9266
20.546 | -103.9211
20.5497 | -103.9144
20.5533 | -103.9076
20.5569 | -103.9009
20.5605 | -103.8944
20.5642 | -103.886
20.5678 | -103.8758
20.5714 | -103.8674
20.5751 | -103.8622
20.5787 | -103.8559
20.5823 | -103.8502
20.5859 | -103.8446
20.5896 | -103.8344
20.5932 | -103.825
20.5968 | -103.8174
20.6005 | -103.8086
20.6041 | -103.7999
20.6077 | -103.7924
20.6113 | -103.7842
20.615 | -103.779
20.6186 | -103.772
20.6222 | -103.7639
20.6258 | -103.7554
20.6295 | -103.7464
20.6331 | -103.7373
20.6367 | -103.7276
20.6404 | -103.7183
20.644 | -103.7093
20.6476 | -103.7016
20.6512 | -103.696
20.6549 | -103.6901
20.6585 | -103.6812
20.6621 | -103.6723
20.6658 | -103.6643
20.6694 | -103.6574
20.673 | -103.6502
20.6766 | -103.6436
20.6803 | -103.6387
20.6839 | -103.635
20.6875 | -103.6299
20.6912 | -103.6258
20.6948 | -103.6216
20.6984 | -103.6192
20.702 | -103.6174
20.7057 | -103.6143
20.7093 | -103.6113
20.7129 | -103.608
20.7166 | -103.6043
20.7202 | -103.6008
20.7238 | -103.5973
20.7274 | -103.5944
20.7311 | -103.594
20.7347 | -103.5941
20.7383 | -103.5923
20.7419 | -103.5913
20.7456 | -103.5896
20.7492 | -103.5887
20.7528 | -103.5883
20.7565 | -103.5872
20.7601 | -103.5866
20.7637 | -103.5847
20.7673 | -103.5815
20.771 | -103.5773
20.7746 | -103.5748
20.7782 | -103.5713
20.7819 | -103.5668
20.7855 | -103.5597
20.7891 | -103.5558
20.7927 | -103.552
20.7964 | -103.5469
20.8 | -103.5418
20.8036 | -103.5391
20.8073 | -103.5363
20.8109 | -103.5341
20.8145 | -103.5292
20.8181 | -103.5235
20.8218 | -103.5191
20.8254 | -103.5163
20.829 | -103.5117
20.8327 | -103.508
20.8363 | -103.5054
20.8399 | -103.503
20.8435 | -103.5023
20.8472 | -103.5031
20.8508 | -103.5041
20.8544 | -103.5043
20.858 | -103.5056
20.8617 | -103.5085
20.8653 | -103.5112
20.8689 | -103.512
20.8726 | -103.5143
20.8762 | -103.5174
20.8798 | -103.5207
20.8834 | -103.5282
20.8871 | -103.5347
20.8907 | -103.5407
20.8943 | -103.5491
20.898 | -103.5579
20.9016 | -103.5678
20.9052 | -103.5772
20.9088 | -103.5846
20.9125 | -103.5922
20.9161 | -103.6027
20.9197 | -103.6135
20.9234 | -103.6244
20.927 | -103.6356
20.9306 | -103.647
20.9342 | -103.6572
20.9379 | -103.668
20.9415 | -103.6805
20.9451 | -103.6919
20.9488 | -103.7025
20.9524 | -103.7152
20.956 | -103.7298
20.9596 | -103.7435
20.9633 | -103.7579
20.9669 | -103.7725
20.9705 | -103.7873
20.9742 | -103.8011
20.9778 | -103.8138
20.9814 | -103.8262
20.985 | -103.84
20.9887 | -103.8541
20.9923 | -103.8683
20.9959 | -103.8817
20.9995 | -103.8979
21.0032 | -103.9147
21.0068 | -103.9298
21.0104 | -103.9459
21.0141 | -103.9613
21.0177 | -103.9766
21.0213 | -103.993
21.0249 | -104.0075
21.0286 | -104.0216
21.0322 | -104.0347
21.0358 | -104.0507
21.0395 | -104.068
21.0431 | -104.0829
21.0467 | -104.0991
21.0503 | -104.1161
21.054 | -104.1335
21.0576 | -104.1508
21.0612 | -104.1696
21.0649 | -104.1904
21.0685 | -104.2126
21.0721 | -104.233
21.0757 | -104.2517
21.0794 | -104.2711
21.083 | -104.2902
21.0866 | -104.3055
21.0902 | -104.3243
21.0939 | -104.3454
21.0975 | -104.3635
21.1011 | -104.38
21.1048 | -104.3963
21.1084 | -104.4109
21.112 | -104.4255
21.1156 | -104.4415
21.1193 | -104.4583
21.1229 | -104.4734
21.1265 | -104.4872
21.1302 | -104.4988
21.1338 | -104.5105
21.1374 | -104.5205
21.141 | -104.5293
21.1447 | -104.5385
21.1483 | -104.545
21.1519 | -104.5509
21.1556 | -104.5567
21.1592 | -104.5609
21.1628 | -104.5629
21.1664 | -104.5667
21.1701 | -104.5713
21.1737 | -104.5735
21.1773 | -104.5739
21.181 | -104.5735
21.1846 | -104.5741
21.1882 | -104.5731
21.1918 | -104.5715
21.1955 | -104.5702
21.1991 | -104.5681
21.2027 | -104.5672
21.2063 | -104.5691
21.21 | -104.5706
21.2136 | -104.5695
21.2172 | -104.5686
21.2209 | -104.5674
21.2245 | -104.5676
21.2281 | -104.5676
21.2317 | -104.5695
21.2354 | -104.5724
21.239 | -104.5735
21.2426 | -104.5739
21.2463 | -104.5758
21.2499 | -104.5774
21.2535 | -104.5784
21.2571 | -104.5799
21.2608 | -104.5814
21.2644 | -104.5835
21.268 | -104.5862
21.2717 | -104.5895
21.2753 | -104.5942
21.2789 | -104.5992
21.2825 | -104.6034
21.2862 | -104.6067
21.2898 | -104.6101
21.2934 | -104.6126
21.2971 | -104.6145
21.3007 | -104.6156
21.3043 | -104.6197
21.3079 | -104.6262
21.3116 | -104.6298
21.3152 | -104.632
21.3188 | -104.6346
21.3224 | -104.6378
21.3261 | -104.6416
21.3297 | -104.644
21.3333 | -104.6448
21.337 | -104.6432
21.3406 | -104.6424
21.3442 | -104.6455
21.3478 | -104.6465
21.3515 | -104.6458
21.3551 | -104.6448
21.3587 | -104.6453
21.3624 | -104.6446
21.366 | -104.6445
21.3696 | -104.6453
21.3732 | -104.6463
21.3769 | -104.6477
21.3805 | -104.6477
21.3841 | -104.6497
21.3878 | -104.6504
21.3914 | -104.6502
21.395 | -104.6471
21.3986 | -104.6463
21.4023 | -104.6451
21.4059 | -104.6433
21.4095 | -104.6423
21.4132 | -104.6417
21.4168 | -104.6363
21.4204 | -104.6294
21.424 | -104.6258
21.4277 | -104.6233
21.4313 | -104.6206
21.4349 | -104.6145
21.4385 | -104.608
21.4422 | -104.6009
21.4458 | -104.5926
21.4494 | -104.5844
21.4531 | -104.5771
21.4567 | -104.5723
21.4603 | -104.5617
21.4639 | -104.5544
21.4676 | -104.5482
21.4712 | -104.5396
21.4748 | -104.53
21.4785 | -104.5227
21.4821 | -104.5149
21.4857 | -104.5047
21.4893 | -104.4957
21.493 | -104.488
21.4966 | -104.4782
21.5002 | -104.4673
21.5039 | -104.4552
21.5075 | -104.4433
21.5111 | -104.4321
21.5147 | -104.42
21.5184 | -104.4075
21.522 | -104.3937
21.5256 | -104.3794
21.5293 | -104.3648
21.5329 | -104.3508
21.5365 | -104.3375
21.5401 | -104.3243
21.5438 | -104.3106
21.5474 | -104.2952
21.551 | -104.2813
21.5546 | -104.268
21.5583 | -104.2547
21.5619 | -104.2434
21.5655 | -104.2318
21.5692 | -104.2185
21.5728 | -104.2054
21.5764 | -104.1926
21.58 | -104.1788
21.5837 | -104.1627
21.5873 | -104.1491
21.5909 | -104.138
21.5946 | -104.127
21.5982 | -104.117
21.6018 | -104.1049
21.6054 | -104.0906
21.6091 | -104.0779
21.6127 | -104.0641
21.6163 | -104.0533
21.62 | -104.0448
21.6236 | -104.0336
21.6272 | -104.0205
21.6308 | -104.0083
21.6345 | -103.9951
21.6381 | -103.9808
21.6417 | -103.9657
21.6454 | -103.9511
21.649 | -103.9376
21.6526 | -103.9256
21.6562 | -103.9156
21.6599 | -103.9051
21.6635 | -103.8928
21.6671 | -103.8804
21.6707 | -103.8695
21.6744 | -103.8586
21.678 | -103.8472
21.6816 | -103.8367
21.6853 | -103.8288
21.6889 | -103.8208
21.6925 | -103.8116
21.6961 | -103.8013
21.6998 | -103.794
21.7034 | -103.7859
21.707 | -103.7785
21.7107 | -103.7723
21.7143 | -103.7667
21.7179 | -103.7585
21.7215 | -103.7544
21.7252 | -103.7495
21.7288 | -103.7405
21.7324 | -103.7341
21.7361 | -103.7301
21.7397 | -103.7277
21.7433 | -103.7249
21.7469 | -103.7221
21.7506 | -103.7188
21.7542 | -103.7155
21.7578 | -103.7117
21.7615 | -103.7045
21.7651 | -103.6992
21.7687 | -103.6973
21.7723 | -103.6934
21.776 | -103.688
21.7796 | -103.6836
21.7832 | -103.6801
21.7868 | -103.6771
21.7905 | -103.6762
21.7941 | -103.6762
21.7977 | -103.6737
21.8014 | -103.669
21.805 | -103.6667
21.8086 | -103.6665
21.8122 | -103.664
21.8159 | -103.6583
21.8195 | -103.6531
21.8231 | -103.6492
21.8268 | -103.6447
21.8304 | -103.6408
21.834 | -103.6372
21.8376 | -103.6319
21.8413 | -103.6251
21.8449 | -103.6183
21.8485 | -103.6122
21.8522 | -103.6049
21.8558 | -103.5984
21.8594 | -103.5903
21.863 | -103.5828
21.8667 | -103.5742
21.8703 | -103.5644
21.8739 | -103.5541
21.8776 | -103.5449
21.8812 | -103.5349
21.8848 | -103.524
21.8884 | -103.5142
21.8921 | -103.5053
21.8957 | -103.4952
21.8993 | -103.4832
21.9029 | -103.4692
21.9066 | -103.4562
21.9102 | -103.4477
21.9138 | -103.4392
21.9175 | -103.4277
21.9211 | -103.4182
21.9247 | -103.4086
21.9283 | -103.3979
21.932 | -103.3893
21.9356 | -103.3816
21.9392 | -103.3728
21.9429 | -103.3657
21.9465 | -103.3617
21.9501 | -103.3583
21.9537 | -103.3534
21.9574 | -103.3468
21.961 | -103.3391
21.9646 | -103.3324
21.9683 | -103.3272
21.9719 | -103.3231
21.9755 | -103.3193
21.9791 | -103.316
21.9828 | -103.3119
21.9864 | -103.3085
21.99 | -103.3068
21.9937 | -103.3052
21.9973 | -103.3035
22.0009 | -103.303
22.0045 | -103.3039
22.0082 | -103.3052
22.0118 | -103.306
22.0154 | -103.3073
22.019 | -103.3087
22.0227 | -103.31
22.0263 | -103.3097
22.0299 | -103.311
22.0336 | -103.3093
22.0372 | -103.3095
22.0408 | -103.3102
22.0444 | -103.3089
22.0481 | -103.3076
22.0517 | -103.3074
22.0553 | -103.3061
22.059 | -103.3053
22.0626 | -103.3055
22.0662 | -103.3057
22.0698 | -103.3057
22.0735 | -103.3066
22.0771 | -103.306
22.0807 | -103.3026
22.0844 | -103.3
22.088 | -103.2978
22.0916 | -103.2951
22.0952 | -103.2923
22.0989 | -103.2894
22.1025 | -103.2853
22.1061 | -103.2786
22.1098 | -103.2743
22.1134 | -103.2727
22.117 | -103.2691
22.1206 | -103.2637
22.1243 | -103.2614
22.1279 | -103.263
22.1315 | -103.2627
22.1351 | -103.2599
22.1388 | -103.2583
22.1424 | -103.2552
22.146 | -103.2514
22.1497 | -103.2477
22.1533 | -103.2441
22.1569 | -103.2406
22.1605 | -103.2377
22.1642 | -103.2375
22.1678 | -103.2358
22.1714 | -103.2327
22.1751 | -103.2295
22.1787 | -103.2268
22.1823 | -103.2233
22.1859 | -103.2193
22.1896 | -103.2148
22.1932 | -103.2124
22.1968 | -103.2084
22.2005 | -103.2045
22.2041 | -103.2023
22.2077 | -103.199
22.2113 | -103.1953
22.215 | -103.1923
22.2186 | -103.1875
22.2222 | -103.1835
22.2259 | -103.1809
22.2295 | -103.177
22.2331 | -103.171
22.2367 | -103.1638
22.2404 | -103.1572
22.244 | -103.151
22.2476 | -103.1438
22.2512 | -103.1382
22.2549 | -103.1296
22.2585 | -103.1206
22.2621 | -103.1128
22.2658 | -103.1024
22.2694 | -103.0913
22.273 | -103.0825
22.2766 | -103.0725
22.2803 | -103.0626
22.2839 | -103.0531
22.2875 | -103.0414
22.2912 | -103.0286
22.2948 | -103.0199
22.2984 | -103.0104
22.302 | -102.9976
22.3057 | -102.9814
22.3093 | -102.964
22.3129 | -102.9475
22.3166 | -102.9318
22.3202 | -102.9157
22.3238 | -102.8994
22.3274 | -102.8822
22.3311 | -102.8645
22.3347 | -102.8482
22.3383 | -102.8321
22.342 | -102.8177
22.3456 | -102.8039
22.3492 | -102.7904
22.3528 | -102.776
22.3565 | -102.7634
22.3601 | -102.7504
22.3637 | -102.7358
22.3673 | -102.7217
22.371 | -102.7104
22.3746 | -102.6981
22.3782 | -102.6855
22.3819 | -102.6716
22.3855 | -102.6579
22.3891 | -102.6441
22.3927 | -102.6328
22.3964 | -102.6227
22.4 | -102.6123
22.4036 | -102.6032
22.4073 | -102.5947
22.4109 | -102.5863
22.4145 | -102.579
22.4181 | -102.5718
22.4218 | -102.5663
22.4254 | -102.5587
22.429 | -102.5514
22.4327 | -102.5457
22.4363 | -102.5415
22.4399 | -102.5368
22.4435 | -102.53
22.4472 | -102.5236
22.4508 | -102.5183
22.4544 | -102.5137
22.4581 | -102.51
22.4617 | -102.5071
22.4653 | -102.5037
22.4689 | -102.4983
22.4726 | -102.4917
22.4762 | -102.4842
22.4798 | -102.4763
22.4834 | -102.4699
22.4871 | -102.4645
22.4907 | -102.4571
22.4943 | -102.4495
22.498 | -102.4428
22.5016 | -102.4363
22.5052 | -102.4251
22.5088 | -102.4204
22.5125 | -102.4122
22.5161 | -102.4046
22.5197 | -102.3962
22.5234 | -102.3854
22.527 | -102.3715
22.5306 | -102.3579
22.5342 | -102.3468
22.5379 | -102.3374
22.5415 | -102.3256
22.5451 | -102.3128
22.5488 | -102.3012
22.5524 | -102.2897
22.556 | -102.2787
22.5596 | -102.2681
22.5633 | -102.2564
22.5669 | -102.2455
22.5705 | -102.2333
22.5741 | -102.2206
22.5778 | -102.2067
22.5814 | -102.1929
22.585 | -102.1812
22.5887 | -102.1693
22.5923 | -102.1548
22.5959 | -102.1403
22.5995 | -102.1256
22.6032 | -102.1106
22.6068 | -102.0986
22.6104 | -102.0871
22.6141 | -102.0754
22.6177 | -102.0643
22.6213 | -102.0542
22.6249 | -102.0474
22.6286 | -102.0405
22.6322 | -102.0338
22.6358 | -102.0267
22.6395 | -102.0204
22.6431 | -102.0134
22.6467 | -102.0057
22.6503 | -101.999
22.654 | -101.9937
22.6576 | -101.9882
22.6612 | -101.9832
22.6649 | -101.9784
22.6685 | -101.9746
22.6721 | -101.9726
22.6757 | -101.9717
22.6794 | -101.9714
22.683 | -101.9708
22.6866 | -101.9707
22.6902 | -101.9712
22.6939 | -101.9732
22.6975 | -101.9761
22.7011 | -101.9811
22.7048 | -101.987
22.7084 | -101.9934
22.712 | -101.9977
22.7156 | -102.002
22.7193 | -102.0087
22.7229 | -102.0141
22.7265 | -102.0199
22.7302 | -102.0252
22.7338 | -102.0292
22.7374 | -102.0362
22.741 | -102.0434
22.7447 | -102.0503
22.7483 | -102.0557
22.7519 | -102.0609
22.7556 | -102.0663
22.7592 | -102.0705
22.7628 | -102.0754
22.7664 | -102.0801
22.7701 | -102.0848
22.7737 | -102.089
22.7773 | -102.0945
22.781 | -102.1005
22.7846 | -102.106
22.7882 | -102.1116
22.7918 | -102.1182
22.7955 | -102.1241
22.7991 | -102.1283
22.8027 | -102.1314
22.8063 | -102.1329
22.81 | -102.1363
22.8136 | -102.1399
22.8172 | -102.1421
22.8209 | -102.1432
22.8245 | -102.1458
22.8281 | -102.1452
22.8317 | -102.1426
22.8354 | -102.1418
22.839 | -102.1418
22.8426 | -102.1429
22.8463 | -102.1463
22.8499 | -102.1508
22.8535 | -102.1532
22.8571 | -102.1524
22.8608 | -102.1478
22.8644 | -102.1421
22.868 | -102.1373
22.8717 | -102.133
22.8753 | -102.1287
22.8789 | -102.1238
22.8825 | -102.1198
22.8862 | -102.1159
22.8898 | -102.1115
22.8934 | -102.1065
22.8971 | -102.1028
22.9007 | -102.0998
22.9043 | -102.0934
22.9079 | -102.0857
22.9116 | -102.0783
22.9152 | -102.0708
22.9188 | -102.0633
22.9224 | -102.0532
22.9261 | -102.0454
22.9297 | -102.0366
22.9333 | -102.0254
22.937 | -102.0118
22.9406 | -102
22.9442 | -101.9908
22.9478 | -101.981
22.9515 | -101.9708
22.9551 | -101.9585
22.9587 | -101.9472
22.9624 | -101.9364
22.966 | -101.9264
22.9696 | -101.9178
22.9732 | -101.9101
22.9769 | -101.9003
22.9805 | -101.8931
22.9841 | -101.8856
22.9878 | -101.8776
22.9914 | -101.8706
22.995 | -101.8639
22.9986 | -101.8569
23.0023 | -101.8504
23.0059 | -101.8451
23.0095 | -101.8393
23.0132 | -101.8347
23.0168 | -101.8288
23.0204 | -101.8244
23.024 | -101.8206
23.0277 | -101.823
23.0313 | -101.8206
23.0349 | -101.8183
23.0385 | -101.8152
23.0422 | -101.813
23.0458 | -101.8102
23.0494 | -101.8087
23.0531 | -101.8069
23.0567 | -101.8066
23.0603 | -101.8074
23.0639 | -101.8066
23.0676 | -101.8075
23.0712 | -101.8084
23.0748 | -101.8105
23.0785 | -101.8121
23.0821 | -101.8121
23.0857 | -101.8123
23.0893 | -101.8104
23.093 | -101.809
23.0966 | -101.8074
23.1002 | -101.8058
23.1039 | -101.8058
23.1075 | -101.8045
23.1111 | -101.8034
23.1147 | -101.8042
23.1184 | -101.8035
23.122 | -101.8026
23.1256 | -101.8028
23.1293 | -101.8033
23.1329 | -101.8025
23.1365 | -101.7988
23.1401 | -101.7954
23.1438 | -101.7916
23.1474 | -101.7882
23.151 | -101.7818
23.1546 | -101.7778
23.1583 | -101.7734
23.1619 | -101.7679
23.1655 | -101.763
23.1692 | -101.7574
23.1728 | -101.7484
23.1764 | -101.7407
23.18 | -101.7329
23.1837 | -101.7271
23.1873 | -101.7226
23.1909 | -101.7173
23.1946 | -101.7119
23.1982 | -101.7051
23.2018 | -101.6976
23.2054 | -101.6914
23.2091 | -101.6847
23.2127 | -101.6779
23.2163 | -101.671
23.22 | -101.6664
23.2236 | -101.6634
23.2272 | -101.6597
23.2308 | -101.6584
23.2345 | -101.6565
23.2381 | -101.6557
23.2417 | -101.6567
23.2454 | -101.6574
23.249 | -101.661
23.2526 | -101.6634
23.2562 | -101.6674
23.2599 | -101.6735
23.2635 | -101.6811
23.2671 | -101.6883
23.2707 | -101.6946
23.2744 | -101.7009
23.278 | -101.7092
23.2816 | -101.7197
23.2853 | -101.73
23.2889 | -101.7421
23.2925 | -101.756
23.2961 | -101.7697
23.2998 | -101.783
23.3034 | -101.7966
23.307 | -101.8093
23.3107 | -101.8226
23.3143 | -101.8387
23.3179 | -101.8544
23.3215 | -101.8691
23.3252 | -101.8833
23.3288 | -101.8975
23.3324 | -101.9141
23.3361 | -101.9318
23.3397 | -101.9472
23.3433 | -101.9639
23.3469 | -101.9797
23.3506 | -101.9954
23.3542 | -102.0103
23.3578 | -102.0231
23.3615 | -102.0352
23.3651 | -102.0501
23.3687 | -102.0666
23.3723 | -102.0812
23.376 | -102.0956
23.3796 | -102.1091
23.3832 | -102.1217
23.3868 | -102.1338
23.3905 | -102.1458
23.3941 | -102.1578
23.3977 | -102.1698
23.4014 | -102.1825
23.405 | -102.1921
23.4086 | -102.2006
23.4122 | -102.2094
23.4159 | -102.2193
23.4195 | -102.2316
23.4231 | -102.2416
23.4268 | -102.2507
23.4304 | -102.2605
23.434 | -102.267
23.4376 | -102.274
23.4413 | -102.2828
23.4449 | -102.2904
23.4485 | -102.2976
23.4522 | -102.3056
23.4558 | -102.3122
23.4594 | -102.3191
23.463 | -102.3285
23.4667 | -102.3379
23.4703 | -102.3448
23.4739 | -102.3541
23.4776 | -102.3643
23.4812 | -102.3754
23.4848 | -102.3848
23.4884 | -102.3948
23.4921 | -102.4051
23.4957 | -102.4127
23.4993 | -102.421
23.5029 | -102.4312
23.5066 | -102.4399
23.5102 | -102.4517
23.5138 | -102.4597
23.5175 | -102.4688
23.5211 | -102.481
23.5247 | -102.493
23.5283 | -102.5053
23.532 | -102.5181
23.5356 | -102.5279
23.5392 | -102.5365
23.5429 | -102.5486
23.5465 | -102.5606
23.5501 | -102.5714
23.5537 | -102.585
23.5574 | -102.5953
23.561 | -102.6065
23.5646 | -102.6188
23.5683 | -102.6321
23.5719 | -102.6422
23.5755 | -102.654
23.5791 | -102.6675
23.5828 | -102.681
23.5864 | -102.6939
23.59 | -102.7052
23.5937 | -102.7174
23.5973 | -102.7319
23.6009 | -102.7473
23.6045 | -102.7614
23.6082 | -102.7764
23.6118 | -102.792
23.6154 | -102.8073
23.619 | -102.8226
23.6227 | -102.8362
23.6263 | -102.848
23.6299 | -102.8617
23.6336 | -102.8745
23.6372 | -102.8896
23.6408 | -102.9049
23.6444 | -102.9205
23.6481 | -102.9348
23.6517 | -102.9499
23.6553 | -102.9647
23.659 | -102.9792
23.6626 | -102.9957
23.6662 | -103.0098
23.6698 | -103.0217
23.6735 | -103.0344
23.6771 | -103.047
23.6807 | -103.0602
23.6844 | -103.073
23.688 | -103.0838
23.6916 | -103.0976
23.6952 | -103.1131
23.6989 | -103.1253
23.7025 | -103.1373
23.7061 | -103.1481
23.7098 | -103.1575
23.7134 | -103.1685
23.717 | -103.1807
23.7206 | -103.1904
23.7243 | -103.1974
23.7279 | -103.2055
23.7315 | -103.2136
23.7351 | -103.2211
23.7388 | -103.2316
23.7424 | -103.2422
23.746 | -103.2525
23.7497 | -103.2616
23.7533 | -103.2687
23.7569 | -103.2762
23.7605 | -103.2823
23.7642 | -103.286
23.7678 | -103.2902
23.7714 | -103.295
23.7751 | -103.2975
23.7787 | -103.2991
23.7823 | -103.3024
23.7859 | -103.3064
23.7896 | -103.3112
23.7932 | -103.3155
23.7968 | -103.3197
23.8005 | -103.3244
23.8041 | -103.3283
23.8077 | -103.3308
23.8113 | -103.332
23.815 | -103.3317
23.8186 | -103.33
23.8222 | -103.3302
23.8258 | -103.3301
23.8295 | -103.3317
23.8331 | -103.3315
23.8367 | -103.333
23.8404 | -103.3365
23.844 | -103.3372
23.8476 | -103.339
23.8512 | -103.34
23.8549 | -103.3407
23.8585 | -103.3406
23.8621 | -103.34
23.8658 | -103.3393
23.8694 | -103.34
23.873 | -103.342
23.8766 | -103.3408
23.8803 | -103.3407
23.8839 | -103.3403
23.8875 | -103.3379
23.8912 | -103.3373
23.8948 | -103.3368
23.8984 | -103.3349
23.902 | -103.3332
23.9057 | -103.3317
23.9093 | -103.3334
23.9129 | -103.3324
23.9166 | -103.331
23.9202 | -103.329
23.9238 | -103.3265
23.9274 | -103.3257
23.9311 | -103.3249
23.9347 | -103.321
23.9383 | -103.3179
23.9419 | -103.3158
23.9456 | -103.3156
23.9492 | -103.3137
23.9528 | -103.309
23.9565 | -103.3068
23.9601 | -103.3047
23.9637 | -103.3026
23.9673 | -103.301
23.971 | -103.2986
23.9746 | -103.2952
23.9782 | -103.2922
23.9819 | -103.2882
23.9855 | -103.2853
23.9891 | -103.2827
23.9927 | -103.2799
23.9964 | -103.2776
24 | -103.2747
24.0036 | -103.2716
24.0073 | -103.2694
24.0109 | -103.2671
24.0145 | -103.2656
24.0181 | -103.2625
24.0218 | -103.2587
24.0254 | -103.2566
24.029 | -103.2528
24.0327 | -103.2474
24.0363 | -103.2436
24.0399 | -103.24
24.0435 | -103.2372
24.0472 | -103.2361
24.0508 | -103.2342
24.0544 | -103.2302
24.0581 | -103.227
24.0617 | -103.2255
24.0653 | -103.2233
24.0689 | -103.2206
24.0726 | -103.218
24.0762 | -103.2155
24.0798 | -103.2131
24.0834 | -103.2107
24.0871 | -103.2084
24.0907 | -103.2063
24.0943 | -103.2055
24.098 | -103.2037
24.1016 | -103.2033
24.1052 | -103.2032
24.1088 | -103.2029
24.1125 | -103.1999
24.1161 | -103.1985
24.1197 | -103.1983
24.1234 | -103.1974
24.127 | -103.1965
24.1306 | -103.1954
24.1342 | -103.1949
24.1379 | -103.1954
24.1415 | -103.1973
24.1451 | -103.2002
24.1488 | -103.201
24.1524 | -103.2027
24.156 | -103.2051
24.1596 | -103.2085
24.1633 | -103.2105
24.1669 | -103.2135
24.1705 | -103.2158
24.1742 | -103.2187
24.1778 | -103.2214
24.1814 | -103.224
24.185 | -103.2275
24.1887 | -103.2312
24.1923 | -103.2325
24.1959 | -103.2333
24.1995 | -103.2355
24.2032 | -103.2389
24.2068 | -103.2434
24.2104 | -103.2477
24.2141 | -103.251
24.2177 | -103.2535
24.2213 | -103.2569
24.2249 | -103.2601
24.2286 | -103.2626
24.2322 | -103.2648
24.2358 | -103.2677
24.2395 | -103.2722
24.2431 | -103.2739
24.2467 | -103.2766
24.2503 | -103.2774
24.254 | -103.2759
24.2576 | -103.275
24.2612 | -103.2754
24.2649 | -103.2768
24.2685 | -103.2764
24.2721 | -103.2738
24.2757 | -103.2711
24.2794 | -103.2683
24.283 | -103.2647
24.2866 | -103.262
24.2902 | -103.2604
24.2939 | -103.261
24.2975 | -103.2595
24.3011 | -103.2586
24.3048 | -103.2572
24.3084 | -103.2538
24.312 | -103.2523
24.3156 | -103.2494
24.3193 | -103.2458
24.3229 | -103.2421
24.3265 | -103.2341
24.3302 | -103.225
24.3338 | -103.219
24.3374 | -103.2148
24.341 | -103.2096
24.3447 | -103.2042
24.3483 | -103.1977
24.3519 | -103.1912
24.3556 | -103.1826
24.3592 | -103.1744
24.3628 | -103.1658
24.3664 | -103.1579
24.3701 | -103.1499
24.3737 | -103.1425
24.3773 | -103.134
24.381 | -103.1261
24.3846 | -103.1197
24.3882 | -103.1119
24.3918 | -103.1034
24.3955 | -103.0973
24.3991 | -103.0918
24.4027 | -103.0874
24.4063 | -103.0819
24.41 | -103.0754
24.4136 | -103.0702
24.4172 | -103.0664
24.4209 | -103.062
24.4245 | -103.0555
24.4281 | -103.0516
24.4317 | -103.0499
24.4354 | -103.0447
24.439 | -103.039
24.4426 | -103.0381
24.4463 | -103.0372
24.4499 | -103.0363
24.4535 | -103.0366
24.4571 | -103.0396
24.4608 | -103.0424
24.4644 | -103.0448
24.468 | -103.0466
24.4717 | -103.0484
24.4753 | -103.0512
24.4789 | -103.0533
24.4825 | -103.0561
24.4862 | -103.0593
24.4898 | -103.0633
24.4934 | -103.0674
24.4971 | -103.0717
24.5007 | -103.0778
24.5043 | -103.0846
24.5079 | -103.0887
24.5116 | -103.0914
24.5152 | -103.0973
24.5188 | -103.1038
24.5224 | -103.1096
24.5261 | -103.113
24.5297 | -103.1163
24.5333 | -103.1202
24.537 | -103.1265
24.5406 | -103.1333
24.5442 | -103.1389
24.5478 | -103.145
24.5515 | -103.15
24.5551 | -103.1527
24.5587 | -103.1523
24.5624 | -103.1494
24.566 | -103.1482
24.5696 | -103.1493
24.5732 | -103.1514
24.5769 | -103.1479
24.5805 | -103.1393
24.5841 | -103.128
24.5878 | -103.1115
24.5914 | -103.0896
24.595 | -103.0401
24.5986 | -102.9159
24.6023 | -102.8643
24.6059 | -102.8175
24.6095 | -102.7715
24.6132 | -102.7223
24.6168 | -102.6703
24.6204 | -102.6124
24.624 | -102.5562
24.6277 | -102.4965
24.6313 | -102.4331
24.6349 | -102.3676
24.6385 | -102.2997
24.6422 | -102.2293
24.6458 | -102.1541
24.6494 | -102.0749
24.6531 | -101.9955
24.6567 | -101.9134
24.6603 | -101.8287
24.6639 | -101.7428
24.6676 | -101.6567
24.6712 | -101.5711
24.6748 | -101.4862
24.6785 | -101.4003
24.6821 | -101.3128
24.6857 | -101.2253
24.6893 | -101.1384
24.693 | -101.0533
24.6966 | -100.9706
24.7002 | -100.8885
24.7039 | -100.8091
24.7075 | -100.7329
24.7111 | -100.6601
24.7147 | -100.5905
24.7184 | -100.5236
24.722 | -100.4598
24.7256 | -100.3992
24.7293 | -100.3428
24.7329 | -100.2892
24.7365 | -100.2384
24.7401 | -100.1929
24.7438 | -100.1535
24.7474 | -100.1193
24.751 | -100.0894
24.7546 | -100.0655
24.7583 | -100.0467
24.7619 | -100.033
24.7655 | -100.024
24.7692 | -100.0209
24.7728 | -100.0236
24.7764 | -100.0307
24.78 | -100.0425
24.7837 | -100.0612
24.7873 | -100.0846
24.7909 | -100.1111
24.7946 | -100.1432
24.7982 | -100.1801
24.8018 | -100.2216
24.8054 | -100.2685
24.8091 | -100.3195
24.8127 | -100.3747
24.8163 | -100.4345
24.82 | -100.498
24.8236 | -100.5655
24.8272 | -100.6366
24.8308 | -100.7102
24.8345 | -100.7861
24.8381 | -100.8653
24.8417 | -100.9484
24.8454 | -101.033
24.849 | -101.118
24.8526 | -101.203
24.8562 | -101.2917
24.8599 | -101.3822
24.8635 | -101.4707
24.8671 | -101.5587
24.8707 | -101.6476
24.8744 | -101.7374
24.878 | -101.8267
24.8816 | -101.9149
24.8853 | -102.0004
24.8889 | -102.0847
24.8925 | -102.167
24.8961 | -102.2463
24.8998 | -102.3232
24.9034 | -102.399
24.907 | -102.4747
24.9107 | -102.5491
24.9143 | -102.6161
24.9179 | -102.679
24.9215 | -102.739
24.9252 | -102.7995
24.9288 | -102.86
24.9324 | -102.9178
24.9361 | -102.9753
24.9397 | -103.0321
24.9433 | -103.1075
24.9469 | -103.1826
24.9506 | -103.247
24.9542 | -103.291
24.9578 | -103.3222
24.9615 | -103.3441
24.9651 | -103.3651
24.9687 | -103.3838
24.9723 | -103.4047
24.976 | -103.4195
24.9796 | -103.4309
24.9832 | -103.4427
24.9868 | -103.4505
24.9905 | -103.4575
24.9941 | -103.4648
24.9977 | -103.4734
25.0014 | -103.4809
25.005 | -103.4863
25.0086 | -103.4915
25.0122 | -103.4974
25.0159 | -103.5055
25.0195 | -103.5118
25.0231 | -103.517
25.0268 | -103.5226
25.0304 | -103.5286
25.034 | -103.5357
25.0376 | -103.5413
25.0413 | -103.5461
25.0449 | -103.5511
25.0485 | -103.5564
25.0522 | -103.5631
25.0558 | -103.5683
25.0594 | -103.5714
25.063 | -103.5748
25.0667 | -103.5761
25.0703 | -103.5786
25.0739 | -103.5822
25.0776 | -103.586
25.0812 | -103.5898
25.0848 | -103.5967
25.0884 | -103.6059
25.0921 | -103.6124
25.0957 | -103.6142
25.0993 | -103.6172
25.1029 | -103.6204
25.1066 | -103.6231
25.1102 | -103.6266
25.1138 | -103.6314
25.1175 | -103.6349
25.1211 | -103.6367
25.1247 | -103.6414
25.1283 | -103.6472
25.132 | -103.6508
25.1356 | -103.654
25.1392 | -103.6569
25.1429 | -103.6621
25.1465 | -103.6683
25.1501 | -103.6729
25.1537 | -103.6786
25.1574 | -103.6826
25.161 | -103.685
25.1646 | -103.6856
25.1683 | -103.6863
25.1719 | -103.6891
25.1755 | -103.6916
25.1791 | -103.6937
25.1828 | -103.6966
25.1864 | -103.6996
25.19 | -103.7037
25.1937 | -103.7055
25.1973 | -103.7055
25.2009 | -103.7066
25.2045 | -103.7092
25.2082 | -103.7116
25.2118 | -103.711
25.2154 | -103.7095
25.219 | -103.7091
25.2227 | -103.7089
25.2263 | -103.7088
25.2299 | -103.7075
25.2336 | -103.7049
25.2372 | -103.703
25.2408 | -103.7022
25.2444 | -103.6991
25.2481 | -103.6958
25.2517 | -103.6927
25.2553 | -103.6907
25.259 | -103.6892
25.2626 | -103.6864
25.2662 | -103.6827
25.2698 | -103.68
25.2735 | -103.6753
25.2771 | -103.6688
25.2807 | -103.6636
25.2844 | -103.6583
25.288 | -103.6518
25.2916 | -103.6448
25.2952 | -103.6393
25.2989 | -103.6335
25.3025 | -103.6297
25.3061 | -103.6268
25.3098 | -103.6224
25.3134 | -103.6165
25.317 | -103.6117
25.3206 | -103.6061
25.3243 | -103.6019
25.3279 | -103.599
25.3315 | -103.5958
25.3351 | -103.5949
25.3388 | -103.594
25.3424 | -103.5937
25.346 | -103.5937
25.3497 | -103.5939
25.3533 | -103.5945
25.3569 | -103.5977
25.3605 | -103.6012
25.3642 | -103.6074
25.3678 | -103.6137
25.3714 | -103.6166
25.3751 | -103.619
25.3787 | -103.623
25.3823 | -103.6289
25.3859 | -103.6346
25.3896 | -103.641
25.3932 | -103.6469
25.3968 | -103.6541
25.4005 | -103.6615
25.4041 | -103.6698
25.4077 | -103.6782
25.4113 | -103.6876
25.415 | -103.6977
25.4186 | -103.7065
25.4222 | -103.7175
25.4259 | -103.7275
25.4295 | -103.7359
25.4331 | -103.744
25.4367 | -103.7537
25.4404 | -103.7651
25.444 | -103.7775
25.4476 | -103.7895
25.4512 | -103.8007
25.4549 | -103.81
25.4585 | -103.82
25.4621 | -103.8308
25.4658 | -103.8416
25.4694 | -103.8522
25.473 | -103.8637
25.4766 | -103.876
25.4803 | -103.8864
25.4839 | -103.8949
25.4875 | -103.9011
25.4912 | -103.9088
25.4948 | -103.9183
25.4984 | -103.928
25.502 | -103.937
25.5057 | -103.9429
25.5093 | -103.9496
25.5129 | -103.9567
25.5166 | -103.9647
25.5202 | -103.9719
25.5238 | -103.981
25.5274 | -103.99
25.5311 | -103.9955
25.5347 | -103.9989
25.5383 | -104.0009
25.5419 | -104.0011
25.5456 | -104.0022
25.5492 | -104.0037
25.5528 | -104.0065
25.5565 | -104.0084
25.5601 | -104.0103
25.5637 | -104.0118
25.5673 | -104.0132
25.571 | -104.013
25.5746 | -104.011
25.5782 | -104.0093
25.5819 | -104.009
25.5855 | -104.0062
25.5891 | -104.0016
25.5927 | -103.9988
25.5964 | -103.9951
25.6 | -103.9907
25.6036 | -103.9876
25.6073 | -103.9836
25.6109 | -103.9801
25.6145 | -103.976
25.6181 | -103.972
25.6218 | -103.9673
25.6254 | -103.9637
25.629 | -103.9595
25.6327 | -103.9572
25.6363 | -103.9555
25.6399 | -103.9501
25.6435 | -103.9456
25.6472 | -103.9427
25.6508 | -103.9394
25.6544 | -103.9371
25.658 | -103.9342
25.6617 | -103.9317
25.6653 | -103.9287
25.6689 | -103.9264
25.6726 | -103.9255
25.6762 | -103.9247
25.6798 | -103.9241
25.6834 | -103.9229
25.6871 | -103.9223
25.6907 | -103.9214
25.6943 | -103.9212
25.698 | -103.9209
25.7016 | -103.9225
25.7052 | -103.9238
25.7088 | -103.9255
25.7125 | -103.9266
25.7161 | -103.9306
25.7197 | -103.9355
25.7234 | -103.9403
25.727 | -103.9444
25.7306 | -103.9475
25.7342 | -103.9537
25.7379 | -103.962
25.7415 | -103.9705
25.7451 | -103.9801
25.7488 | -103.9917
25.7524 | -104.0021
25.756 | -104.0142
25.7596 | -104.0268
25.7633 | -104.0386
25.7669 | -104.0492
25.7705 | -104.0615
25.7741 | -104.0744
25.7778 | -104.0885
25.7814 | -104.1029
25.785 | -104.1169
25.7887 | -104.133
25.7923 | -104.1496
25.7959 | -104.1676
25.7995 | -104.1821
25.8032 | -104.1951
25.8068 | -104.21
25.8104 | -104.2267
25.8141 | -104.2446
25.8177 | -104.2616
25.8213 | -104.2771
25.8249 | -104.2916
25.8286 | -104.3082
25.8322 | -104.3245
25.8358 | -104.3409
25.8395 | -104.3569
25.8431 | -104.3727
25.8467 | -104.3893
25.8503 | -104.4083
25.854 | -104.4257
25.8576 | -104.4408
25.8612 | -104.4551
25.8649 | -104.4687
25.8685 | -104.4836
25.8721 | -104.4975
25.8757 | -104.5108
25.8794 | -104.5228
25.883 | -104.5364
25.8866 | -104.549
25.8902 | -104.5602
25.8939 | -104.5724
25.8975 | -104.5845
25.9011 | -104.5961
25.9048 | -104.6068
25.9084 | -104.6162
25.912 | -104.6257
25.9156 | -104.6351
25.9193 | -104.6448
25.9229 | -104.6521
25.9265 | -104.6588
25.9302 | -104.6655
25.9338 | -104.6728
25.9374 | -104.6797
25.941 | -104.6863
25.9447 | -104.6916
25.9483 | -104.6975
25.9519 | -104.7047
25.9556 | -104.7091
25.9592 | -104.7134
25.9628 | -104.7161
25.9664 | -104.7169
25.9701 | -104.719
25.9737 | -104.7214
25.9773 | -104.7238
25.981 | -104.7265
25.9846 | -104.7288
25.9882 | -104.7309
25.9918 | -104.7341
25.9955 | -104.7347
25.9991 | -104.7369
26.0027 | -104.7422
26.0063 | -104.7442
26.01 | -104.7448
26.0136 | -104.7452
26.0172 | -104.7466
26.0209 | -104.749
26.0245 | -104.7505
26.0281 | -104.7508
26.0317 | -104.7509
26.0354 | -104.7522
26.039 | -104.7547
26.0426 | -104.7578
26.0463 | -104.7612
26.0499 | -104.7663
26.0535 | -104.7705
26.0571 | -104.7741
26.0608 | -104.7773
26.0644 | -104.7809
26.068 | -104.7864
26.0717 | -104.7927
26.0753 | -104.7982
26.0789 | -104.8051
26.0825 | -104.8111
26.0862 | -104.8165
26.0898 | -104.8222
26.0934 | -104.8283
26.0971 | -104.836
26.1007 | -104.8428
26.1043 | -104.8494
26.1079 | -104.855
26.1116 | -104.8598
26.1152 | -104.8659
26.1188 | -104.8731
26.1224 | -104.8815
26.1261 | -104.8914
26.1297 | -104.9015
26.1333 | -104.9122
26.137 | -104.9221
26.1406 | -104.9307
26.1442 | -104.9401
26.1478 | -104.9493
26.1515 | -104.9589
26.1551 | -104.968
26.1587 | -104.9764
26.1624 | -104.9847
26.166 | -104.9955
26.1696 | -105.0027
26.1732 | -105.0115
26.1769 | -105.019
26.1805 | -105.0257
26.1841 | -105.0319
26.1878 | -105.039
26.1914 | -105.0476
26.195 | -105.0554
26.1986 | -105.0591
26.2023 | -105.0631
26.2059 | -105.0682
26.2095 | -105.0748
26.2132 | -105.0808
26.2168 | -105.086
26.2204 | -105.0885
26.224 | -105.0916
26.2277 | -105.096
26.2313 | -105.1006
26.2349 | -105.1044
26.2385 | -105.1081
26.2422 | -105.1105
26.2458 | -105.1131
26.2494 | -105.116
26.2531 | -105.1161
26.2567 | -105.1173
26.2603 | -105.1193
26.2639 | -105.1201
26.2676 | -105.1219
26.2712 | -105.1221
26.2748 | -105.1225
26.2785 | -105.1222
26.2821 | -105.1191
26.2857 | -105.1144
26.2893 | -105.1105
26.293 | -105.1087
26.2966 | -105.1071
26.3002 | -105.1065
26.3039 | -105.1053
26.3075 | -105.1032
26.3111 | -105.1006
26.3147 | -105.0971
26.3184 | -105.0935
26.322 | -105.092
26.3256 | -105.0903
26.3293 | -105.0887
26.3329 | -105.0878
26.3365 | -105.0858
26.3401 | -105.0858
26.3438 | -105.0842
26.3474 | -105.0812
26.351 | -105.0782
26.3546 | -105.0777
26.3583 | -105.0761
26.3619 | -105.0735
26.3655 | -105.0715
26.3692 | -105.069
26.3728 | -105.0661
26.3764 | -105.0645
26.38 | -105.0659
26.3837 | -105.0679
26.3873 | -105.0692
26.3909 | -105.0696
26.3946 | -105.0701
26.3982 | -105.0717
26.4018 | -105.0737
26.4054 | -105.0728
26.4091 | -105.0688
26.4127 | -105.0667
26.4163 | -105.0674
26.42 | -105.0682
26.4236 | -105.0702
26.4272 | -105.0717
26.4308 | -105.0737
26.4345 | -105.0747
26.4381 | -105.072
26.4417 | -105.0711
26.4454 | -105.0738
26.449 | -105.0762
26.4526 | -105.0764
26.4562 | -105.0783
26.4599 | -105.0805
26.4635 | -105.0806
26.4671 | -105.0798
26.4707 | -105.0801
26.4744 | -105.0811
26.478 | -105.0835
26.4816 | -105.085
26.4853 | -105.0849
26.4889 | -105.0838
26.4925 | -105.0817
26.4961 | -105.0779
26.4998 | -105.0745
26.5034 | -105.0706
26.507 | -105.0676
26.5107 | -105.0662
26.5143 | -105.0615
26.5179 | -105.059
26.5215 | -105.0568
26.5252 | -105.0532
26.5288 | -105.05
26.5324 | -105.0462
26.5361 | -105.0409
26.5397 | -105.0344
26.5433 | -105.029
26.5469 | -105.0243
26.5506 | -105.0186
26.5542 | -105.0114
26.5578 | -105.0063
26.5615 | -105.0009
26.5651 | -104.9931
26.5687 | -104.9858
26.5723 | -104.9798
26.576 | -104.9719
26.5796 | -104.963
26.5832 | -104.9515
26.5868 | -104.9414
26.5905 | -104.9306
26.5941 | -104.9206
26.5977 | -104.9101
26.6014 | -104.8971
26.605 | -104.8841
26.6086 | -104.872
26.6122 | -104.8614
26.6159 | -104.853
26.6195 | -104.8433
26.6231 | -104.832
26.6268 | -104.8211
26.6304 | -104.8095
26.634 | -104.7977
26.6376 | -104.7862
26.6413 | -104.7758
26.6449 | -104.7665
26.6485 | -104.7572
26.6522 | -104.7485
26.6558 | -104.7397
26.6594 | -104.7321
26.663 | -104.7244
26.6667 | -104.7163
26.6703 | -104.7086
26.6739 | -104.7017
26.6776 | -104.6949
26.6812 | -104.6888
26.6848 | -104.6823
26.6884 | -104.675
26.6921 | -104.6731
26.6957 | -104.6671
26.6993 | -104.6615
26.7029 | -104.6576
26.7066 | -104.6532
26.7102 | -104.6496
26.7138 | -104.6477
26.7175 | -104.6438
26.7211 | -104.6402
26.7247 | -104.638
26.7283 | -104.6361
26.732 | -104.6333
26.7356 | -104.6325
26.7392 | -104.6314
26.7429 | -104.6314
26.7465 | -104.6313
26.7501 | -104.6318
26.7537 | -104.6326
26.7574 | -104.6314
26.761 | -104.629
26.7646 | -104.6268
26.7683 | -104.6267
26.7719 | -104.6264
26.7755 | -104.6256
26.7791 | -104.627
26.7828 | -104.6286
26.7864 | -104.6321
26.79 | -104.6327
26.7937 | -104.632
26.7973 | -104.6326
26.8009 | -104.6316
26.8045 | -104.6329
26.8082 | -104.6341
26.8118 | -104.6374
26.8154 | -104.6421
26.819 | -104.6453
26.8227 | -104.6467
26.8263 | -104.6474
26.8299 | -104.6499
26.8336 | -104.6521
26.8372 | -104.6545
26.8408 | -104.6586
26.8444 | -104.6627
26.8481 | -104.6661
26.8517 | -104.6697
26.8553 | -104.6717
26.859 | -104.6732
26.8626 | -104.6737
26.8662 | -104.6744
26.8698 | -104.6768
26.8735 | -104.6806
26.8771 | -104.6839
26.8807 | -104.6867
26.8844 | -104.6881
26.888 | -104.6874
26.8916 | -104.6891
26.8952 | -104.6897
26.8989 | -104.6901
26.9025 | -104.691
26.9061 | -104.6922
26.9098 | -104.6936
26.9134 | -104.6944
26.917 | -104.696
26.9206 | -104.6962
26.9243 | -104.6945
26.9279 | -104.6942
26.9315 | -104.6948
26.9351 | -104.6927
26.9388 | -104.6909
26.9424 | -104.6892
26.946 | -104.6894
26.9497 | -104.6898
26.9533 | -104.6912
26.9569 | -104.6928
26.9605 | -104.6948
26.9642 | -104.6957
26.9678 | -104.6977
26.9714 | -104.6989
26.9751 | -104.7002
26.9787 | -104.7031
26.9823 | -104.7054
26.9859 | -104.7072
26.9896 | -104.7079
26.9932 | -104.7071
26.9968 | -104.7085
27.0005 | -104.7111
27.0041 | -104.7132
27.0077 | -104.7153
27.0113 | -104.7166
27.015 | -104.7168
27.0186 | -104.7162
27.0222 | -104.7161
27.0258 | -104.7174
27.0295 | -104.7175
27.0331 | -104.7174
27.0367 | -104.7175
27.0404 | -104.7182
27.044 | -104.7186
27.0476 | -104.718
27.0512 | -104.7171
27.0549 | -104.7172
27.0585 | -104.7179
27.0621 | -104.7189
27.0658 | -104.7189
27.0694 | -104.7173
27.073 | -104.716
27.0766 | -104.7151
27.0803 | -104.7137
27.0839 | -104.7129
27.0875 | -104.7124
27.0912 | -104.713
27.0948 | -104.7128
27.0984 | -104.7113
27.102 | -104.7116
27.1057 | -104.7119
27.1093 | -104.7117
27.1129 | -104.7116
27.1166 | -104.7113
27.1202 | -104.7134
27.1238 | -104.7149
27.1274 | -104.7161
27.1311 | -104.7167
27.1347 | -104.717
27.1383 | -104.7168
27.142 | -104.72
27.1456 | -104.7241
27.1492 | -104.7255
27.1528 | -104.7269
27.1565 | -104.7301
27.1601 | -104.7314
27.1637 | -104.7342
27.1673 | -104.7375
27.171 | -104.7421
27.1746 | -104.7461
27.1782 | -104.7484
27.1819 | -104.7517
27.1855 | -104.7534
27.1891 | -104.7555
27.1927 | -104.756
27.1964 | -104.7566
27.2 | -104.7586
27.2036 | -104.7604
27.2073 | -104.7628
27.2109 | -104.7647
27.2145 | -104.7665
27.2181 | -104.7673
27.2218 | -104.7682
27.2254 | -104.7687
27.229 | -104.7684
27.2327 | -104.7673
27.2363 | -104.7661
27.2399 | -104.7665
27.2435 | -104.7678
27.2472 | -104.7697
27.2508 | -104.7709
27.2544 | -104.7706
27.2581 | -104.7701
27.2617 | -104.7684
27.2653 | -104.7676
27.2689 | -104.7678
27.2726 | -104.7692
27.2762 | -104.7688
27.2798 | -104.7662
27.2834 | -104.7626
27.2871 | -104.76
27.2907 | -104.757
27.2943 | -104.7533
27.298 | -104.7503
27.3016 | -104.7479
27.3052 | -104.7468
27.3088 | -104.7457
27.3125 | -104.743
27.3161 | -104.7411
27.3197 | -104.7406
27.3234 | -104.74
27.327 | -104.7388
27.3306 | -104.7343
27.3342 | -104.7309
27.3379 | -104.726
27.3415 | -104.7208
27.3451 | -104.7171
27.3488 | -104.7141
27.3524 | -104.7108
27.356 | -104.7089
27.3596 | -104.7062
27.3633 | -104.7018
27.3669 | -104.6969
27.3705 | -104.6935
27.3742 | -104.6913
27.3778 | -104.6898
27.3814 | -104.686
27.385 | -104.6831
27.3887 | -104.6816
27.3923 | -104.6792
27.3959 | -104.6768
27.3995 | -104.677
27.4032 | -104.6772
27.4068 | -104.6768
27.4104 | -104.6761
27.4141 | -104.6744
27.4177 | -104.6728
27.4213 | -104.6719
27.4249 | -104.6689
27.4286 | -104.667
27.4322 | -104.6671
27.4358 | -104.6662
27.4395 | -104.6655
27.4431 | -104.6651
27.4467 | -104.6646
27.4503 | -104.6635
27.454 | -104.662
27.4576 | -104.6605
27.4612 | -104.6597
27.4649 | -104.6582
27.4685 | -104.6538
27.4721 | -104.6512
27.4757 | -104.6484
27.4794 | -104.6477
27.483 | -104.649
27.4866 | -104.6475
27.4902 | -104.6467
27.4939 | -104.6449
27.4975 | -104.6437
27.5011 | -104.644
27.5048 | -104.6427
27.5084 | -104.6416
27.512 | -104.6423
27.5156 | -104.643
27.5193 | -104.6442
27.5229 | -104.6451
27.5265 | -104.6432
27.5302 | -104.6411
27.5338 | -104.6411
27.5374 | -104.6421
27.541 | -104.6423
27.5447 | -104.6417
27.5483 | -104.6397
27.5519 | -104.6363
27.5556 | -104.6328
27.5592 | -104.6297
27.5628 | -104.6263
27.5664 | -104.6234
27.5701 | -104.6211
27.5737 | -104.6177
27.5773 | -104.6137
27.581 | -104.6086
27.5846 | -104.6039
27.5882 | -104.5979
27.5918 | -104.5935
27.5955 | -104.59
27.5991 | -104.5853
27.6027 | -104.5812
27.6063 | -104.5764
27.61 | -104.5704
27.6136 | -104.5658
27.6172 | -104.5614
27.6209 | -104.5578
27.6245 | -104.5534
27.6281 | -104.5464
27.6317 | -104.5391
27.6354 | -104.534
27.639 | -104.5307
27.6426 | -104.5248
27.6463 | -104.5189
27.6499 | -104.5144
27.6535 | -104.5092
27.6571 | -104.5036
27.6608 | -104.4955
27.6644 | -104.4879
27.668 | -104.4818
27.6717 | -104.4772
27.6753 | -104.4722
27.6789 | -104.4668
27.6825 | -104.4637
27.6862 | -104.4593
27.6898 | -104.4531
27.6934 | -104.4476
27.6971 | -104.4427
27.7007 | -104.4387
27.7043 | -104.4342
27.7079 | -104.4287
27.7116 | -104.4224
27.7152 | -104.4168
27.7188 | -104.4131
27.7224 | -104.4085
27.7261 | -104.406
27.7297 | -104.4021
27.7333 | -104.3973
27.737 | -104.3913
27.7406 | -104.3843
27.7442 | -104.3794
27.7478 | -104.3741
27.7515 | -104.3696
27.7551 | -104.3664
27.7587 | -104.3619
27.7624 | -104.3547
27.766 | -104.3468
27.7696 | -104.3378
27.7732 | -104.3291
27.7769 | -104.3217
27.7805 | -104.3149
27.7841 | -104.308
27.7878 | -104.2993
27.7914 | -104.2912
27.795 | -104.2841
27.7986 | -104.2765
27.8023 | -104.2685
27.8059 | -104.2625
27.8095 | -104.2572
27.8132 | -104.2501
27.8168 | -104.2434
27.8204 | -104.2363
27.824 | -104.2277
27.8277 | -104.2209
27.8313 | -104.2154
27.8349 | -104.2091
27.8385 | -104.2034
27.8422 | -104.1989
27.8458 | -104.1959
27.8494 | -104.193
27.8531 | -104.1888
27.8567 | -104.1852
27.8603 | -104.1831
27.8639 | -104.1797
27.8676 | -104.1769
27.8712 | -104.1733
27.8748 | -104.1691
27.8785 | -104.1672
27.8821 | -104.1648
27.8857 | -104.1629
27.8893 | -104.1596
27.893 | -104.1533
27.8966 | -104.148
27.9002 | -104.1428
27.9039 | -104.1402
27.9075 | -104.1392
27.9111 | -104.1383
27.9147 | -104.1391
27.9184 | -104.1393
27.922 | -104.1367
27.9256 | -104.1313
27.9293 | -104.1265
27.9329 | -104.1245
27.9365 | -104.1242
27.9401 | -104.1255
27.9438 | -104.1276
27.9474 | -104.1282
27.951 | -104.1258
27.9546 | -104.1256
27.9583 | -104.1251
27.9619 | -104.1246
27.9655 | -104.1256
27.9692 | -104.126
27.9728 | -104.1257
27.9764 | -104.1247
27.98 | -104.124
27.9837 | -104.1232
27.9873 | -104.1197
27.9909 | -104.1169
27.9946 | -104.1141
27.9982 | -104.1102
28.0018 | -104.1064