## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 237 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 230
Hood | 02-92B/P | 169
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1003 | hPa |  | PRC PRO 01
Humidity | 55.6 | %/Hr
Internal temperature of calibration bench | 23.29 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1011 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.2 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.95 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.06 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3060 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 11/06/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 12/06/2018
Technical Manager | Anthony HUE | Approval |  | 13/06/2018
CEO | Robert GLEMOT | Approval |  | 13/06/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3492 | 57.5169
0.012209 | -35.9313 | 38.3201
0.018313 | -34.8805 | 28.1806
0.024417 | -34.4164 | 21.9438
0.030521 | -34.3015 | 18.1639
0.036626 | -34.1568 | 15.108
0.04273 | -34.0618 | 12.7589
0.048834 | -34.0159 | 11.2417
0.054938 | -33.9527 | 10.1024
0.061043 | -33.8941 | 9.0263
0.067147 | -33.8827 | 8.2076
0.073251 | -33.8821 | 7.4505
0.079355 | -33.8988 | 6.7916
0.08546 | -33.9122 | 6.201
0.091564 | -33.9079 | 5.7576
0.097668 | -33.8907 | 5.4974
0.10377 | -33.8692 | 5.2435
0.10988 | -33.8526 | 4.9618
0.11598 | -33.8443 | 4.6673
0.12209 | -33.8469 | 4.3836
0.12819 | -33.8571 | 4.1031
0.13429 | -33.8672 | 3.8201
0.1404 | -33.8683 | 3.5525
0.1465 | -33.8685 | 3.3047
0.15261 | -33.8642 | 3.0932
0.15871 | -33.857 | 2.9352
0.16482 | -33.8489 | 2.8306
0.17092 | -33.8477 | 2.7827
0.17702 | -33.8464 | 2.7056
0.18313 | -33.8466 | 2.6243
0.18923 | -33.8472 | 2.5161
0.19534 | -33.8468 | 2.393
0.20144 | -33.8462 | 2.2575
0.20754 | -33.8436 | 2.1424
0.21365 | -33.843 | 2.0321
0.21975 | -33.8427 | 1.9373
0.22586 | -33.844 | 1.8512
0.23196 | -33.8466 | 1.7783
0.23807 | -33.8507 | 1.7133
0.24417 | -33.8549 | 1.6524
0.25027 | -33.8589 | 1.5883
0.25638 | -33.8614 | 1.5247
0.26248 | -33.8632 | 1.4576
0.26859 | -33.8642 | 1.3913
0.27469 | -33.8644 | 1.3198
0.2808 | -33.8652 | 1.246
0.2869 | -33.8658 | 1.1747
0.293 | -33.8665 | 1.1072
0.29911 | -33.867 | 1.0464
0.30521 | -33.8677 | 0.98709
0.31132 | -33.8682 | 0.94735
0.31742 | -33.8691 | 0.90366
0.32353 | -33.8692 | 0.86757
0.32963 | -33.8699 | 0.83687
0.33573 | -33.8708 | 0.81475
0.34184 | -33.872 | 0.79164
0.34794 | -33.8741 | 0.76613
0.35405 | -33.8762 | 0.74001
0.36015 | -33.8786 | 0.71266
0.36626 | -33.8806 | 0.68354
0.37236 | -33.8827 | 0.64514
0.37846 | -33.885 | 0.60009
0.38457 | -33.886 | 0.54349
0.39067 | -33.8871 | 0.4967
0.39678 | -33.8878 | 0.44579
0.40288 | -33.8884 | 0.39349
0.40899 | -33.8883 | 0.34333
0.41509 | -33.8877 | 0.2948
0.42119 | -33.8876 | 0.25525
0.4273 | -33.8877 | 0.21208
0.4334 | -33.8876 | 0.16982
0.43951 | -33.8878 | 0.12983
0.44561 | -33.8885 | 0.093139
0.45172 | -33.89 | 0.06317
0.45782 | -33.8907 | 0.028851
0.46392 | -33.8919 | -0.0053447
0.47003 | -33.8929 | -0.036159
0.47613 | -33.8937 | -0.069359
0.48224 | -33.8942 | -0.10338
0.48834 | -33.8934 | -0.13864
0.49445 | -33.8931 | -0.17413
0.50055 | -33.8926 | -0.20685
0.50665 | -33.8919 | -0.24153
0.51276 | -33.8909 | -0.2765
0.51886 | -33.8898 | -0.31118
0.52497 | -33.8886 | -0.33604
0.53107 | -33.8878 | -0.36604
0.53717 | -33.8871 | -0.39469
0.54328 | -33.8863 | -0.41786
0.54938 | -33.886 | -0.44087
0.55549 | -33.886 | -0.46238
0.56159 | -33.887 | -0.47413
0.5677 | -33.8873 | -0.48889
0.5738 | -33.8883 | -0.50558
0.5799 | -33.889 | -0.52096
0.58601 | -33.89 | -0.53159
0.59211 | -33.8916 | -0.53854
0.59822 | -33.8929 | -0.5485
0.60432 | -33.8942 | -0.56103
0.61043 | -33.8954 | -0.56964
0.61653 | -33.8966 | -0.58078
0.62263 | -33.898 | -0.59085
0.62874 | -33.8986 | -0.60273
0.63484 | -33.8994 | -0.61309
0.64095 | -33.8999 | -0.62424
0.64705 | -33.9002 | -0.63504
0.65316 | -33.9003 | -0.64834
0.65926 | -33.9003 | -0.66462
0.66536 | -33.8992 | -0.68238
0.67147 | -33.8987 | -0.69666
0.67757 | -33.8983 | -0.71171
0.68368 | -33.8974 | -0.73084
0.68978 | -33.8966 | -0.74788
0.69589 | -33.8955 | -0.76726
0.70199 | -33.8945 | -0.78882
0.70809 | -33.8936 | -0.80795
0.7142 | -33.8928 | -0.82833
0.7203 | -33.8916 | -0.84669
0.72641 | -33.8909 | -0.86833
0.73251 | -33.8902 | -0.8876
0.73862 | -33.8902 | -0.9081
0.74472 | -33.8899 | -0.92718
0.75082 | -33.8896 | -0.94839
0.75693 | -33.8898 | -0.96913
0.76303 | -33.8902 | -0.98985
0.76914 | -33.892 | -1.0096
0.77524 | -33.8929 | -1.0283
0.78135 | -33.8941 | -1.0476
0.78745 | -33.8955 | -1.0669
0.79355 | -33.8967 | -1.0856
0.79966 | -33.8984 | -1.1046
0.80576 | -33.9008 | -1.1241
0.81187 | -33.9025 | -1.1451
0.81797 | -33.9045 | -1.1656
0.82408 | -33.9063 | -1.1856
0.83018 | -33.9082 | -1.206
0.83628 | -33.9102 | -1.2242
0.84239 | -33.9116 | -1.245
0.84849 | -33.9132 | -1.2665
0.8546 | -33.9148 | -1.2863
0.8607 | -33.9163 | -1.307
0.86681 | -33.9176 | -1.3259
0.87291 | -33.9188 | -1.3446
0.87901 | -33.9194 | -1.3631
0.88512 | -33.9204 | -1.3814
0.89122 | -33.9213 | -1.3998
0.89733 | -33.9219 | -1.416
0.90343 | -33.9225 | -1.4337
0.90953 | -33.9228 | -1.4484
0.91564 | -33.9229 | -1.4646
0.92174 | -33.9233 | -1.4801
0.92785 | -33.9235 | -1.4949
0.93395 | -33.9238 | -1.509
0.94006 | -33.924 | -1.5225
0.94616 | -33.9237 | -1.5352
0.95226 | -33.9239 | -1.5484
0.95837 | -33.924 | -1.5628
0.96447 | -33.9242 | -1.5762
0.97058 | -33.9241 | -1.5902
0.97668 | -33.9239 | -1.6042
0.98279 | -33.9236 | -1.6176
0.98889 | -33.9234 | -1.6291
0.99499 | -33.9234 | -1.6434
1.0011 | -33.9232 | -1.6547
1.0072 | -33.9228 | -1.6673
1.0133 | -33.9225 | -1.6792
1.0194 | -33.9221 | -1.693
1.0255 | -33.9217 | -1.7074
1.0316 | -33.9212 | -1.7201
1.0377 | -33.9208 | -1.7347
1.0438 | -33.9203 | -1.748
1.0499 | -33.9195 | -1.7639
1.056 | -33.9189 | -1.7791
1.0621 | -33.9183 | -1.7939
1.0682 | -33.9179 | -1.8073
1.0743 | -33.9172 | -1.8233
1.0805 | -33.9167 | -1.8382
1.0866 | -33.9158 | -1.854
1.0927 | -33.915 | -1.8691
1.0988 | -33.9144 | -1.8847
1.1049 | -33.9136 | -1.8992
1.111 | -33.9128 | -1.915
1.1171 | -33.9121 | -1.93
1.1232 | -33.9116 | -1.9453
1.1293 | -33.9112 | -1.9597
1.1354 | -33.9109 | -1.9749
1.1415 | -33.9106 | -1.9908
1.1476 | -33.91 | -2.0063
1.1537 | -33.9095 | -2.0214
1.1598 | -33.9096 | -2.0362
1.1659 | -33.9093 | -2.0497
1.172 | -33.9091 | -2.064
1.1781 | -33.9089 | -2.0781
1.1842 | -33.9087 | -2.0923
1.1903 | -33.9091 | -2.1028
1.1964 | -33.9091 | -2.116
1.2025 | -33.9092 | -2.1299
1.2086 | -33.9095 | -2.1423
1.2147 | -33.9096 | -2.1547
1.2209 | -33.9099 | -2.1658
1.227 | -33.9107 | -2.1757
1.2331 | -33.9112 | -2.1876
1.2392 | -33.9115 | -2.1992
1.2453 | -33.912 | -2.2099
1.2514 | -33.9125 | -2.2197
1.2575 | -33.9131 | -2.2301
1.2636 | -33.914 | -2.2382
1.2697 | -33.9146 | -2.2467
1.2758 | -33.9153 | -2.2575
1.2819 | -33.9161 | -2.2667
1.288 | -33.9168 | -2.2766
1.2941 | -33.9177 | -2.2846
1.3002 | -33.9184 | -2.2919
1.3063 | -33.919 | -2.3003
1.3124 | -33.9196 | -2.3087
1.3185 | -33.9204 | -2.3171
1.3246 | -33.9211 | -2.3254
1.3307 | -33.9219 | -2.3333
1.3368 | -33.9227 | -2.3412
1.3429 | -33.9233 | -2.3496
1.349 | -33.9238 | -2.3582
1.3551 | -33.9245 | -2.3668
1.3613 | -33.925 | -2.3745
1.3674 | -33.9256 | -2.3838
1.3735 | -33.926 | -2.393
1.3796 | -33.9264 | -2.4014
1.3857 | -33.9267 | -2.4099
1.3918 | -33.9272 | -2.4195
1.3979 | -33.9277 | -2.4281
1.404 | -33.9281 | -2.4392
1.4101 | -33.9284 | -2.4491
1.4162 | -33.9287 | -2.4595
1.4223 | -33.9291 | -2.4703
1.4284 | -33.9295 | -2.48
1.4345 | -33.9297 | -2.4909
1.4406 | -33.9299 | -2.5045
1.4467 | -33.9302 | -2.5157
1.4528 | -33.9306 | -2.5269
1.4589 | -33.9309 | -2.5388
1.465 | -33.9311 | -2.5492
1.4711 | -33.9315 | -2.5602
1.4772 | -33.9318 | -2.5734
1.4833 | -33.9321 | -2.5861
1.4894 | -33.9325 | -2.5983
1.4955 | -33.9329 | -2.6102
1.5016 | -33.9333 | -2.6224
1.5078 | -33.9336 | -2.6361
1.5139 | -33.9339 | -2.649
1.52 | -33.9341 | -2.6615
1.5261 | -33.9343 | -2.6743
1.5322 | -33.9346 | -2.6882
1.5383 | -33.9349 | -2.7007
1.5444 | -33.9353 | -2.7132
1.5505 | -33.9356 | -2.7263
1.5566 | -33.9361 | -2.7393
1.5627 | -33.9364 | -2.7528
1.5688 | -33.9369 | -2.7665
1.5749 | -33.9371 | -2.7791
1.581 | -33.9374 | -2.7924
1.5871 | -33.9377 | -2.8061
1.5932 | -33.938 | -2.8187
1.5993 | -33.9383 | -2.8329
1.6054 | -33.9386 | -2.8461
1.6115 | -33.939 | -2.86
1.6176 | -33.9393 | -2.8724
1.6237 | -33.9394 | -2.8858
1.6298 | -33.9396 | -2.8983
1.6359 | -33.94 | -2.9109
1.642 | -33.9403 | -2.9228
1.6482 | -33.9406 | -2.9355
1.6543 | -33.9409 | -2.9489
1.6604 | -33.9412 | -2.9607
1.6665 | -33.9413 | -2.9731
1.6726 | -33.9416 | -2.9855
1.6787 | -33.9419 | -2.9978
1.6848 | -33.9419 | -3.0101
1.6909 | -33.9422 | -3.0228
1.697 | -33.9424 | -3.035
1.7031 | -33.9425 | -3.0476
1.7092 | -33.9426 | -3.0596
1.7153 | -33.9428 | -3.0716
1.7214 | -33.9427 | -3.0841
1.7275 | -33.9427 | -3.0965
1.7336 | -33.9427 | -3.1091
1.7397 | -33.943 | -3.1215
1.7458 | -33.943 | -3.1336
1.7519 | -33.9431 | -3.1457
1.758 | -33.9432 | -3.158
1.7641 | -33.9435 | -3.1698
1.7702 | -33.9435 | -3.1827
1.7763 | -33.9435 | -3.1944
1.7824 | -33.9436 | -3.2071
1.7885 | -33.9436 | -3.2197
1.7947 | -33.9438 | -3.2316
1.8008 | -33.9439 | -3.2431
1.8069 | -33.944 | -3.255
1.813 | -33.9441 | -3.2676
1.8191 | -33.9441 | -3.2805
1.8252 | -33.9442 | -3.2938
1.8313 | -33.9444 | -3.3061
1.8374 | -33.9446 | -3.3184
1.8435 | -33.9447 | -3.3306
1.8496 | -33.9448 | -3.344
1.8557 | -33.945 | -3.357
1.8618 | -33.9453 | -3.3697
1.8679 | -33.9454 | -3.3819
1.874 | -33.9455 | -3.3942
1.8801 | -33.9457 | -3.4069
1.8862 | -33.946 | -3.4192
1.8923 | -33.9461 | -3.4317
1.8984 | -33.9462 | -3.4441
1.9045 | -33.9463 | -3.4561
1.9106 | -33.9464 | -3.4687
1.9167 | -33.9465 | -3.4812
1.9228 | -33.9466 | -3.4933
1.9289 | -33.9467 | -3.5056
1.9351 | -33.9469 | -3.5169
1.9412 | -33.9471 | -3.5285
1.9473 | -33.9474 | -3.5408
1.9534 | -33.9476 | -3.5526
1.9595 | -33.9478 | -3.5648
1.9656 | -33.948 | -3.5773
1.9717 | -33.9482 | -3.5893
1.9778 | -33.9485 | -3.601
1.9839 | -33.9488 | -3.6128
1.99 | -33.9489 | -3.624
1.9961 | -33.949 | -3.6364
2.0022 | -33.9492 | -3.648
2.0083 | -33.9493 | -3.6594
2.0144 | -33.9495 | -3.6709
2.0205 | -33.9498 | -3.6828
2.0266 | -33.9501 | -3.6942
2.0327 | -33.9504 | -3.706
2.0388 | -33.9507 | -3.7177
2.0449 | -33.9509 | -3.7291
2.051 | -33.9511 | -3.7407
2.0571 | -33.9514 | -3.752
2.0632 | -33.9516 | -3.7633
2.0693 | -33.9518 | -3.7756
2.0754 | -33.952 | -3.7868
2.0816 | -33.9523 | -3.7987
2.0877 | -33.9525 | -3.81
2.0938 | -33.9529 | -3.8223
2.0999 | -33.953 | -3.8341
2.106 | -33.9532 | -3.8458
2.1121 | -33.9534 | -3.8577
2.1182 | -33.9536 | -3.8685
2.1243 | -33.9539 | -3.8792
2.1304 | -33.9541 | -3.8902
2.1365 | -33.9543 | -3.9019
2.1426 | -33.9545 | -3.9134
2.1487 | -33.9548 | -3.9247
2.1548 | -33.9552 | -3.9363
2.1609 | -33.9556 | -3.9477
2.167 | -33.956 | -3.9587
2.1731 | -33.9563 | -3.9704
2.1792 | -33.9565 | -3.9812
2.1853 | -33.9569 | -3.9927
2.1914 | -33.9571 | -4.0039
2.1975 | -33.9575 | -4.0154
2.2036 | -33.9577 | -4.0274
2.2097 | -33.9579 | -4.0387
2.2158 | -33.9583 | -4.0492
2.222 | -33.9585 | -4.0596
2.2281 | -33.9587 | -4.0707
2.2342 | -33.9591 | -4.082
2.2403 | -33.9592 | -4.0924
2.2464 | -33.9595 | -4.1031
2.2525 | -33.9598 | -4.1143
2.2586 | -33.9601 | -4.1247
2.2647 | -33.9604 | -4.135
2.2708 | -33.9607 | -4.1461
2.2769 | -33.961 | -4.1567
2.283 | -33.9612 | -4.1678
2.2891 | -33.9614 | -4.1785
2.2952 | -33.9616 | -4.1889
2.3013 | -33.9618 | -4.1997
2.3074 | -33.9621 | -4.21
2.3135 | -33.9623 | -4.2212
2.3196 | -33.9625 | -4.2314
2.3257 | -33.9627 | -4.2421
2.3318 | -33.9629 | -4.2531
2.3379 | -33.9631 | -4.2637
2.344 | -33.9632 | -4.2749
2.3501 | -33.9634 | -4.2856
2.3562 | -33.9636 | -4.296
2.3623 | -33.9637 | -4.3063
2.3685 | -33.9638 | -4.3167
2.3746 | -33.964 | -4.3279
2.3807 | -33.9642 | -4.3382
2.3868 | -33.9644 | -4.349
2.3929 | -33.9646 | -4.3596
2.399 | -33.9646 | -4.37
2.4051 | -33.9649 | -4.3806
2.4112 | -33.9651 | -4.3911
2.4173 | -33.9653 | -4.4017
2.4234 | -33.9654 | -4.4124
2.4295 | -33.9656 | -4.4231
2.4356 | -33.9657 | -4.434
2.4417 | -33.9658 | -4.4451
2.4478 | -33.9659 | -4.4558
2.4539 | -33.9662 | -4.4659
2.46 | -33.9663 | -4.4774
2.4661 | -33.9664 | -4.4883
2.4722 | -33.9666 | -4.4987
2.4783 | -33.9667 | -4.5092
2.4844 | -33.9668 | -4.5201
2.4905 | -33.9669 | -4.5313
2.4966 | -33.9671 | -4.5423
2.5027 | -33.9672 | -4.5529
2.5089 | -33.9673 | -4.5636
2.515 | -33.9674 | -4.5744
2.5211 | -33.9675 | -4.5853
2.5272 | -33.9676 | -4.5963
2.5333 | -33.9677 | -4.6073
2.5394 | -33.9679 | -4.618
2.5455 | -33.968 | -4.6293
2.5516 | -33.9682 | -4.6399
2.5577 | -33.9684 | -4.6515
2.5638 | -33.9686 | -4.6633
2.5699 | -33.9686 | -4.6749
2.576 | -33.9687 | -4.6864
2.5821 | -33.9688 | -4.6977
2.5882 | -33.969 | -4.7087
2.5943 | -33.9691 | -4.7197
2.6004 | -33.9694 | -4.7309
2.6065 | -33.9695 | -4.7413
2.6126 | -33.9696 | -4.7518
2.6187 | -33.9698 | -4.7624
2.6248 | -33.9699 | -4.7733
2.6309 | -33.9701 | -4.7848
2.637 | -33.9704 | -4.7965
2.6431 | -33.9705 | -4.807
2.6492 | -33.9707 | -4.8185
2.6554 | -33.9709 | -4.8296
2.6615 | -33.971 | -4.8405
2.6676 | -33.9713 | -4.852
2.6737 | -33.9715 | -4.8633
2.6798 | -33.9718 | -4.8745
2.6859 | -33.9721 | -4.8859
2.692 | -33.9722 | -4.8969
2.6981 | -33.9724 | -4.9081
2.7042 | -33.9727 | -4.9193
2.7103 | -33.9729 | -4.9301
2.7164 | -33.9732 | -4.9411
2.7225 | -33.9734 | -4.9526
2.7286 | -33.9737 | -4.9637
2.7347 | -33.9739 | -4.9743
2.7408 | -33.9741 | -4.9847
2.7469 | -33.9744 | -4.996
2.753 | -33.9747 | -5.0068
2.7591 | -33.975 | -5.0182
2.7652 | -33.9752 | -5.0298
2.7713 | -33.9754 | -5.0407
2.7774 | -33.9757 | -5.0517
2.7835 | -33.9758 | -5.0629
2.7896 | -33.976 | -5.0735
2.7958 | -33.9763 | -5.0844
2.8019 | -33.9766 | -5.095
2.808 | -33.9769 | -5.1059
2.8141 | -33.9772 | -5.1173
2.8202 | -33.9775 | -5.1283
2.8263 | -33.9778 | -5.1393
2.8324 | -33.9781 | -5.1509
2.8385 | -33.9785 | -5.1619
2.8446 | -33.9788 | -5.1729
2.8507 | -33.9792 | -5.1839
2.8568 | -33.9795 | -5.1945
2.8629 | -33.9798 | -5.2052
2.869 | -33.9801 | -5.2166
2.8751 | -33.9804 | -5.2274
2.8812 | -33.9808 | -5.2389
2.8873 | -33.9811 | -5.25
2.8934 | -33.9814 | -5.2612
2.8995 | -33.9817 | -5.2724
2.9056 | -33.982 | -5.2836
2.9117 | -33.9823 | -5.2948
2.9178 | -33.9826 | -5.3057
2.9239 | -33.983 | -5.3174
2.93 | -33.9834 | -5.329
2.9361 | -33.9837 | -5.34
2.9423 | -33.984 | -5.3509
2.9484 | -33.9844 | -5.3622
2.9545 | -33.9847 | -5.3734
2.9606 | -33.9851 | -5.385
2.9667 | -33.9855 | -5.3962
2.9728 | -33.9858 | -5.4077
2.9789 | -33.9862 | -5.419
2.985 | -33.9866 | -5.4304
2.9911 | -33.9869 | -5.4418
2.9972 | -33.9874 | -5.453
3.0033 | -33.9877 | -5.4645
3.0094 | -33.9881 | -5.4757
3.0155 | -33.9885 | -5.4868
3.0216 | -33.9888 | -5.4982
3.0277 | -33.9891 | -5.5099
3.0338 | -33.9895 | -5.5215
3.0399 | -33.9899 | -5.5328
3.046 | -33.9902 | -5.5446
3.0521 | -33.9905 | -5.5559
3.0582 | -33.9908 | -5.5679
3.0643 | -33.9911 | -5.5796
3.0704 | -33.9914 | -5.5914
3.0765 | -33.9916 | -5.6026
3.0827 | -33.992 | -5.614
3.0888 | -33.9923 | -5.6253
3.0949 | -33.9927 | -5.6366
3.101 | -33.993 | -5.6481
3.1071 | -33.9933 | -5.6598
3.1132 | -33.9937 | -5.6716
3.1193 | -33.994 | -5.6832
3.1254 | -33.9944 | -5.6945
3.1315 | -33.9947 | -5.7064
3.1376 | -33.995 | -5.7177
3.1437 | -33.9953 | -5.7292
3.1498 | -33.9957 | -5.7412
3.1559 | -33.996 | -5.7527
3.162 | -33.9962 | -5.7648
3.1681 | -33.9965 | -5.7767
3.1742 | -33.9968 | -5.7883
3.1803 | -33.9971 | -5.7999
3.1864 | -33.9975 | -5.8112
3.1925 | -33.9978 | -5.8227
3.1986 | -33.9981 | -5.8345
3.2047 | -33.9983 | -5.8465
3.2108 | -33.9986 | -5.8575
3.2169 | -33.9989 | -5.8692
3.223 | -33.9991 | -5.881
3.2292 | -33.9994 | -5.8927
3.2353 | -33.9996 | -5.9035
3.2414 | -33.9999 | -5.9149
3.2475 | -34.0001 | -5.9263
3.2536 | -34.0004 | -5.9376
3.2597 | -34.0006 | -5.9486
3.2658 | -34.0009 | -5.9595
3.2719 | -34.0011 | -5.9704
3.278 | -34.0014 | -5.9819
3.2841 | -34.0016 | -5.993
3.2902 | -34.0019 | -6.0043
3.2963 | -34.0021 | -6.0157
3.3024 | -34.0023 | -6.0267
3.3085 | -34.0026 | -6.0382
3.3146 | -34.0027 | -6.0493
3.3207 | -34.003 | -6.0606
3.3268 | -34.0033 | -6.072
3.3329 | -34.0035 | -6.0827
3.339 | -34.0037 | -6.0935
3.3451 | -34.0039 | -6.1047
3.3512 | -34.0041 | -6.1153
3.3573 | -34.0043 | -6.1265
3.3634 | -34.0045 | -6.1376
3.3696 | -34.0047 | -6.1489
3.3757 | -34.0049 | -6.1596
3.3818 | -34.0052 | -6.1705
3.3879 | -34.0054 | -6.1815
3.394 | -34.0055 | -6.1929
3.4001 | -34.0057 | -6.2036
3.4062 | -34.0059 | -6.2145
3.4123 | -34.0061 | -6.2251
3.4184 | -34.0063 | -6.2363
3.4245 | -34.0065 | -6.2471
3.4306 | -34.0066 | -6.2578
3.4367 | -34.0068 | -6.2684
3.4428 | -34.007 | -6.2788
3.4489 | -34.0072 | -6.2897
3.455 | -34.0074 | -6.3005
3.4611 | -34.0076 | -6.3117
3.4672 | -34.0078 | -6.3228
3.4733 | -34.008 | -6.3338
3.4794 | -34.0082 | -6.3446
3.4855 | -34.0083 | -6.355
3.4916 | -34.0086 | -6.3654
3.4977 | -34.0087 | -6.3767
3.5038 | -34.0089 | -6.3876
3.5099 | -34.0091 | -6.3979
3.5161 | -34.0092 | -6.4084
3.5222 | -34.0094 | -6.4188
3.5283 | -34.0095 | -6.4296
3.5344 | -34.0097 | -6.4404
3.5405 | -34.0099 | -6.4509
3.5466 | -34.0101 | -6.4618
3.5527 | -34.0102 | -6.4727
3.5588 | -34.0103 | -6.4834
3.5649 | -34.0104 | -6.494
3.571 | -34.0106 | -6.5047
3.5771 | -34.0108 | -6.5158
3.5832 | -34.0109 | -6.5259
3.5893 | -34.0112 | -6.5365
3.5954 | -34.0114 | -6.5474
3.6015 | -34.0115 | -6.5583
3.6076 | -34.0118 | -6.5689
3.6137 | -34.0119 | -6.579
3.6198 | -34.0122 | -6.5894
3.6259 | -34.0124 | -6.6004
3.632 | -34.0126 | -6.6111
3.6381 | -34.0128 | -6.6219
3.6442 | -34.0129 | -6.6325
3.6503 | -34.0132 | -6.6428
3.6565 | -34.0134 | -6.6537
3.6626 | -34.0135 | -6.6644
3.6687 | -34.0138 | -6.6752
3.6748 | -34.0141 | -6.6858
3.6809 | -34.0143 | -6.6965
3.687 | -34.0145 | -6.7075
3.6931 | -34.0148 | -6.7179
3.6992 | -34.015 | -6.7286
3.7053 | -34.0152 | -6.7392
3.7114 | -34.0154 | -6.75
3.7175 | -34.0156 | -6.7608
3.7236 | -34.0159 | -6.771
3.7297 | -34.0161 | -6.7819
3.7358 | -34.0163 | -6.7929
3.7419 | -34.0165 | -6.8038
3.748 | -34.0168 | -6.8145
3.7541 | -34.017 | -6.8252
3.7602 | -34.0173 | -6.8358
3.7663 | -34.0175 | -6.8464
3.7724 | -34.0178 | -6.8572
3.7785 | -34.018 | -6.8677
3.7846 | -34.0182 | -6.8782
3.7907 | -34.0186 | -6.8893
3.7969 | -34.0188 | -6.9001
3.803 | -34.0191 | -6.9104
3.8091 | -34.0193 | -6.9211
3.8152 | -34.0196 | -6.9316
3.8213 | -34.0198 | -6.9425
3.8274 | -34.0201 | -6.953
3.8335 | -34.0204 | -6.9638
3.8396 | -34.0207 | -6.9745
3.8457 | -34.021 | -6.9853
3.8518 | -34.0214 | -6.996
3.8579 | -34.0217 | -7.0066
3.864 | -34.022 | -7.0175
3.8701 | -34.0223 | -7.0285
3.8762 | -34.0225 | -7.0391
3.8823 | -34.0228 | -7.05
3.8884 | -34.0231 | -7.0607
3.8945 | -34.0235 | -7.0717
3.9006 | -34.0238 | -7.0827
3.9067 | -34.0241 | -7.0938
3.9128 | -34.0244 | -7.1046
3.9189 | -34.0247 | -7.1152
3.925 | -34.0249 | -7.126
3.9311 | -34.0252 | -7.1362
3.9372 | -34.0255 | -7.1471
3.9434 | -34.0258 | -7.1579
3.9495 | -34.0261 | -7.1687
3.9556 | -34.0264 | -7.1793
3.9617 | -34.0267 | -7.1897
3.9678 | -34.0271 | -7.2001
3.9739 | -34.0273 | -7.2107
3.98 | -34.0276 | -7.2216
3.9861 | -34.0279 | -7.2322
3.9922 | -34.0281 | -7.2429
3.9983 | -34.0285 | -7.2533
4.0044 | -34.0287 | -7.2635
4.0105 | -34.029 | -7.2741
4.0166 | -34.0293 | -7.2848
4.0227 | -34.0296 | -7.2955
4.0288 | -34.0299 | -7.3066
4.0349 | -34.0301 | -7.317
4.041 | -34.0304 | -7.3276
4.0471 | -34.0306 | -7.3383
4.0532 | -34.0309 | -7.3486
4.0593 | -34.0312 | -7.3594
4.0654 | -34.0315 | -7.3698
4.0715 | -34.0318 | -7.3809
4.0776 | -34.0321 | -7.3914
4.0838 | -34.0324 | -7.4025
4.0899 | -34.0327 | -7.4128
4.096 | -34.033 | -7.4234
4.1021 | -34.0332 | -7.4339
4.1082 | -34.0335 | -7.4442
4.1143 | -34.0337 | -7.4545
4.1204 | -34.034 | -7.4648
4.1265 | -34.0343 | -7.4751
4.1326 | -34.0346 | -7.4854
4.1387 | -34.0349 | -7.4958
4.1448 | -34.0352 | -7.5063
4.1509 | -34.0354 | -7.5168
4.157 | -34.0357 | -7.5278
4.1631 | -34.0359 | -7.5386
4.1692 | -34.0362 | -7.5489
4.1753 | -34.0365 | -7.5594
4.1814 | -34.0367 | -7.5694
4.1875 | -34.037 | -7.5802
4.1936 | -34.0373 | -7.5909
4.1997 | -34.0375 | -7.6011
4.2058 | -34.0377 | -7.6118
4.2119 | -34.038 | -7.6221
4.218 | -34.0382 | -7.6322
4.2241 | -34.0384 | -7.6425
4.2303 | -34.0387 | -7.6528
4.2364 | -34.039 | -7.6633
4.2425 | -34.0392 | -7.6737
4.2486 | -34.0394 | -7.6843
4.2547 | -34.0395 | -7.6945
4.2608 | -34.0398 | -7.705
4.2669 | -34.04 | -7.7151
4.273 | -34.0402 | -7.7257
4.2791 | -34.0405 | -7.7361
4.2852 | -34.0407 | -7.7466
4.2913 | -34.0408 | -7.7568
4.2974 | -34.0411 | -7.7674
4.3035 | -34.0413 | -7.778
4.3096 | -34.0415 | -7.7886
4.3157 | -34.0418 | -7.7988
4.3218 | -34.042 | -7.8091
4.3279 | -34.0422 | -7.8195
4.334 | -34.0424 | -7.8298
4.3401 | -34.0426 | -7.8398
4.3462 | -34.0429 | -7.8502
4.3523 | -34.0432 | -7.8602
4.3584 | -34.0433 | -7.8701
4.3645 | -34.0435 | -7.8804
4.3707 | -34.0437 | -7.8904
4.3768 | -34.044 | -7.9007
4.3829 | -34.0442 | -7.9114
4.389 | -34.0444 | -7.9216
4.3951 | -34.0446 | -7.9317
4.4012 | -34.0448 | -7.9423
4.4073 | -34.045 | -7.9525
4.4134 | -34.0452 | -7.9627
4.4195 | -34.0455 | -7.9729
4.4256 | -34.0457 | -7.9831
4.4317 | -34.046 | -7.9933
4.4378 | -34.0461 | -8.0032
4.4439 | -34.0463 | -8.0132
4.45 | -34.0465 | -8.0234
4.4561 | -34.0467 | -8.0336
4.4622 | -34.0469 | -8.044
4.4683 | -34.0471 | -8.0541
4.4744 | -34.0473 | -8.0647
4.4805 | -34.0475 | -8.0745
4.4866 | -34.0477 | -8.0847
4.4927 | -34.0479 | -8.0951
4.4988 | -34.0482 | -8.1053
4.5049 | -34.0484 | -8.1151
4.511 | -34.0486 | -8.1251
4.5172 | -34.0489 | -8.1352
4.5233 | -34.049 | -8.1457
4.5294 | -34.0493 | -8.1558
4.5355 | -34.0495 | -8.1661
4.5416 | -34.0497 | -8.1762
4.5477 | -34.0499 | -8.1864
4.5538 | -34.0501 | -8.1962
4.5599 | -34.0503 | -8.2063
4.566 | -34.0505 | -8.2166
4.5721 | -34.0508 | -8.2268
4.5782 | -34.051 | -8.2372
4.5843 | -34.0512 | -8.2475
4.5904 | -34.0515 | -8.2577
4.5965 | -34.0517 | -8.2678
4.6026 | -34.0519 | -8.2784
4.6087 | -34.0521 | -8.2885
4.6148 | -34.0523 | -8.299
4.6209 | -34.0526 | -8.3091
4.627 | -34.0528 | -8.3196
4.6331 | -34.053 | -8.3297
4.6392 | -34.0532 | -8.34
4.6453 | -34.0534 | -8.3502
4.6514 | -34.0537 | -8.3605
4.6576 | -34.0539 | -8.3711
4.6637 | -34.054 | -8.3809
4.6698 | -34.0542 | -8.391
4.6759 | -34.0545 | -8.4016
4.682 | -34.0547 | -8.4117
4.6881 | -34.0548 | -8.4219
4.6942 | -34.055 | -8.4321
4.7003 | -34.0553 | -8.4424
4.7064 | -34.0555 | -8.4524
4.7125 | -34.0557 | -8.4629
4.7186 | -34.0559 | -8.4729
4.7247 | -34.0562 | -8.4832
4.7308 | -34.0564 | -8.4934
4.7369 | -34.0566 | -8.5037
4.743 | -34.0568 | -8.5143
4.7491 | -34.0571 | -8.5249
4.7552 | -34.0573 | -8.5352
4.7613 | -34.0575 | -8.5456
4.7674 | -34.0577 | -8.5557
4.7735 | -34.0579 | -8.5664
4.7796 | -34.0581 | -8.5769
4.7857 | -34.0584 | -8.5873
4.7918 | -34.0586 | -8.5974
4.7979 | -34.0589 | -8.6076
4.8041 | -34.0591 | -8.6179
4.8102 | -34.0593 | -8.6279
4.8163 | -34.0595 | -8.6381
4.8224 | -34.0598 | -8.6485
4.8285 | -34.06 | -8.6584
4.8346 | -34.0603 | -8.6685
4.8407 | -34.0604 | -8.6789
4.8468 | -34.0607 | -8.6895
4.8529 | -34.061 | -8.7
4.859 | -34.0612 | -8.7106
4.8651 | -34.0614 | -8.7207
4.8712 | -34.0617 | -8.7308
4.8773 | -34.0618 | -8.7412
4.8834 | -34.0621 | -8.7515
4.8895 | -34.0623 | -8.7616
4.8956 | -34.0625 | -8.7719
4.9017 | -34.0627 | -8.782
4.9078 | -34.0629 | -8.7926
4.9139 | -34.0632 | -8.8026
4.92 | -34.0635 | -8.8126
4.9261 | -34.0637 | -8.8226
4.9322 | -34.064 | -8.833
4.9383 | -34.0642 | -8.8432
4.9445 | -34.0645 | -8.8535
4.9506 | -34.0647 | -8.8637
4.9567 | -34.065 | -8.8737
4.9628 | -34.0652 | -8.884
4.9689 | -34.0654 | -8.8946
4.975 | -34.0656 | -8.9052
4.9811 | -34.0659 | -8.9152
4.9872 | -34.0662 | -8.9254
4.9933 | -34.0665 | -8.936
4.9994 | -34.0667 | -8.9466
5.0055 | -34.067 | -8.957
5.0116 | -34.0673 | -8.9677
5.0177 | -34.0676 | -8.9783
5.0238 | -34.0678 | -8.9889
5.0299 | -34.0682 | -8.9996
5.036 | -34.0684 | -9.0099
5.0421 | -34.0686 | -9.0202
5.0482 | -34.0689 | -9.0305
5.0543 | -34.0692 | -9.0406
5.0604 | -34.0695 | -9.051
5.0665 | -34.0698 | -9.0611
5.0726 | -34.07 | -9.0717
5.0787 | -34.0703 | -9.0823
5.0848 | -34.0706 | -9.0927
5.091 | -34.0708 | -9.1029
5.0971 | -34.0711 | -9.1129
5.1032 | -34.0714 | -9.1236
5.1093 | -34.0717 | -9.134
5.1154 | -34.0719 | -9.1441
5.1215 | -34.0721 | -9.1543
5.1276 | -34.0724 | -9.1644
5.1337 | -34.0727 | -9.1745
5.1398 | -34.073 | -9.1846
5.1459 | -34.0733 | -9.1952
5.152 | -34.0735 | -9.2058
5.1581 | -34.0738 | -9.2163
5.1642 | -34.0741 | -9.2267
5.1703 | -34.0744 | -9.2371
5.1764 | -34.0747 | -9.2473
5.1825 | -34.075 | -9.2574
5.1886 | -34.0752 | -9.2675
5.1947 | -34.0755 | -9.2779
5.2008 | -34.0758 | -9.2884
5.2069 | -34.0761 | -9.299
5.213 | -34.0764 | -9.3092
5.2191 | -34.0767 | -9.3196
5.2252 | -34.0769 | -9.33
5.2314 | -34.0772 | -9.3406
5.2375 | -34.0775 | -9.3507
5.2436 | -34.0777 | -9.3612
5.2497 | -34.078 | -9.3713
5.2558 | -34.0783 | -9.382
5.2619 | -34.0786 | -9.3927
5.268 | -34.0789 | -9.4032
5.2741 | -34.0792 | -9.4136
5.2802 | -34.0795 | -9.4244
5.2863 | -34.0798 | -9.4347
5.2924 | -34.0801 | -9.4453
5.2985 | -34.0805 | -9.4557
5.3046 | -34.0807 | -9.4662
5.3107 | -34.081 | -9.4768
5.3168 | -34.0813 | -9.4873
5.3229 | -34.0816 | -9.4978
5.329 | -34.0819 | -9.5082
5.3351 | -34.0822 | -9.5183
5.3412 | -34.0825 | -9.5287
5.3473 | -34.0828 | -9.5392
5.3534 | -34.0831 | -9.5498
5.3595 | -34.0834 | -9.5602
5.3656 | -34.0837 | -9.5707
5.3717 | -34.084 | -9.5815
5.3779 | -34.0843 | -9.5922
5.384 | -34.0846 | -9.6026
5.3901 | -34.0849 | -9.6133
5.3962 | -34.0852 | -9.6236
5.4023 | -34.0855 | -9.6343
5.4084 | -34.0858 | -9.6447
5.4145 | -34.0862 | -9.6549
5.4206 | -34.0864 | -9.6654
5.4267 | -34.0867 | -9.6762
5.4328 | -34.087 | -9.6867
5.4389 | -34.0873 | -9.697
5.445 | -34.0876 | -9.7078
5.4511 | -34.0879 | -9.718
5.4572 | -34.0883 | -9.7285
5.4633 | -34.0886 | -9.739
5.4694 | -34.0888 | -9.7495
5.4755 | -34.0891 | -9.7598
5.4816 | -34.0894 | -9.7703
5.4877 | -34.0898 | -9.7806
5.4938 | -34.0901 | -9.7908
5.4999 | -34.0904 | -9.8012
5.506 | -34.0907 | -9.8116
5.5121 | -34.091 | -9.822
5.5183 | -34.0913 | -9.8323
5.5244 | -34.0916 | -9.843
5.5305 | -34.0919 | -9.8534
5.5366 | -34.0922 | -9.8636
5.5427 | -34.0925 | -9.8742
5.5488 | -34.0928 | -9.8847
5.5549 | -34.0931 | -9.8952
5.561 | -34.0934 | -9.9059
5.5671 | -34.0937 | -9.9162
5.5732 | -34.094 | -9.9269
5.5793 | -34.0943 | -9.9377
5.5854 | -34.0947 | -9.9481
5.5915 | -34.095 | -9.9589
5.5976 | -34.0953 | -9.9695
5.6037 | -34.0956 | -9.98
5.6098 | -34.0959 | -9.9904
5.6159 | -34.0962 | -10.0008
5.622 | -34.0965 | -10.0113
5.6281 | -34.0968 | -10.0219
5.6342 | -34.0971 | -10.0327
5.6403 | -34.0974 | -10.0432
5.6464 | -34.0977 | -10.0538
5.6525 | -34.098 | -10.0644
5.6586 | -34.0984 | -10.0746
5.6648 | -34.0987 | -10.0851
5.6709 | -34.099 | -10.0958
5.677 | -34.0993 | -10.1063
5.6831 | -34.0997 | -10.1169
5.6892 | -34.0999 | -10.1276
5.6953 | -34.1002 | -10.138
5.7014 | -34.1006 | -10.1487
5.7075 | -34.1008 | -10.1591
5.7136 | -34.1011 | -10.1696
5.7197 | -34.1014 | -10.1801
5.7258 | -34.1017 | -10.1904
5.7319 | -34.1021 | -10.2009
5.738 | -34.1024 | -10.2114
5.7441 | -34.1027 | -10.2219
5.7502 | -34.1031 | -10.2321
5.7563 | -34.1034 | -10.2426
5.7624 | -34.1037 | -10.2533
5.7685 | -34.104 | -10.2638
5.7746 | -34.1043 | -10.2746
5.7807 | -34.1046 | -10.2851
5.7868 | -34.1049 | -10.2958
5.7929 | -34.1052 | -10.3062
5.799 | -34.1055 | -10.3166
5.8052 | -34.1058 | -10.3269
5.8113 | -34.1061 | -10.3373
5.8174 | -34.1064 | -10.3477
5.8235 | -34.1067 | -10.3581
5.8296 | -34.107 | -10.3688
5.8357 | -34.1073 | -10.3792
5.8418 | -34.1076 | -10.3898
5.8479 | -34.1079 | -10.4004
5.854 | -34.1082 | -10.411
5.8601 | -34.1085 | -10.4214
5.8662 | -34.1088 | -10.432
5.8723 | -34.1091 | -10.4423
5.8784 | -34.1094 | -10.4527
5.8845 | -34.1097 | -10.463
5.8906 | -34.11 | -10.4736
5.8967 | -34.1103 | -10.4839
5.9028 | -34.1106 | -10.4942
5.9089 | -34.111 | -10.5045
5.915 | -34.1113 | -10.5148
5.9211 | -34.1116 | -10.5252
5.9272 | -34.1119 | -10.5354
5.9333 | -34.1122 | -10.5457
5.9394 | -34.1125 | -10.5562
5.9455 | -34.1128 | -10.5664
5.9517 | -34.1132 | -10.5767
5.9578 | -34.1135 | -10.5872
5.9639 | -34.1138 | -10.5973
5.97 | -34.1142 | -10.6078
5.9761 | -34.1145 | -10.6181
5.9822 | -34.1148 | -10.6286
5.9883 | -34.1151 | -10.639
5.9944 | -34.1154 | -10.6495
6.0005 | -34.1158 | -10.6601
6.0066 | -34.1161 | -10.6703
6.0127 | -34.1164 | -10.6804
6.0188 | -34.1167 | -10.6911
6.0249 | -34.117 | -10.7016
6.031 | -34.1173 | -10.7118
6.0371 | -34.1176 | -10.7222
6.0432 | -34.1179 | -10.7325
6.0493 | -34.1183 | -10.7427
6.0554 | -34.1186 | -10.7529
6.0615 | -34.1189 | -10.7633
6.0676 | -34.1193 | -10.7739
6.0737 | -34.1196 | -10.7844
6.0798 | -34.1199 | -10.7947
6.0859 | -34.1202 | -10.805
6.0921 | -34.1205 | -10.8152
6.0982 | -34.1209 | -10.8256
6.1043 | -34.1212 | -10.8358
6.1104 | -34.1216 | -10.8464
6.1165 | -34.1219 | -10.8563
6.1226 | -34.1222 | -10.8665
6.1287 | -34.1225 | -10.8768
6.1348 | -34.1229 | -10.8872
6.1409 | -34.1232 | -10.8974
6.147 | -34.1235 | -10.9077
6.1531 | -34.1238 | -10.918
6.1592 | -34.1241 | -10.9282
6.1653 | -34.1244 | -10.9384
6.1714 | -34.1247 | -10.9487
6.1775 | -34.1251 | -10.9591
6.1836 | -34.1254 | -10.9695
6.1897 | -34.1257 | -10.9797
6.1958 | -34.126 | -10.9903
6.2019 | -34.1264 | -11.0007
6.208 | -34.1267 | -11.011
6.2141 | -34.127 | -11.0213
6.2202 | -34.1274 | -11.0317
6.2263 | -34.1277 | -11.0419
6.2325 | -34.128 | -11.0518
6.2386 | -34.1283 | -11.0618
6.2447 | -34.1286 | -11.0724
6.2508 | -34.129 | -11.0826
6.2569 | -34.1293 | -11.0928
6.263 | -34.1296 | -11.1031
6.2691 | -34.1299 | -11.1133
6.2752 | -34.1302 | -11.1236
6.2813 | -34.1306 | -11.1339
6.2874 | -34.1309 | -11.144
6.2935 | -34.1312 | -11.1542
6.2996 | -34.1315 | -11.1645
6.3057 | -34.1318 | -11.1748
6.3118 | -34.1321 | -11.1851
6.3179 | -34.1324 | -11.1953
6.324 | -34.1327 | -11.2055
6.3301 | -34.1331 | -11.2157
6.3362 | -34.1334 | -11.2259
6.3423 | -34.1337 | -11.2363
6.3484 | -34.1341 | -11.2466
6.3545 | -34.1344 | -11.2565
6.3606 | -34.1347 | -11.2665
6.3667 | -34.1349 | -11.2765
6.3728 | -34.1353 | -11.2868
6.379 | -34.1356 | -11.297
6.3851 | -34.1359 | -11.307
6.3912 | -34.1362 | -11.3172
6.3973 | -34.1365 | -11.3274
6.4034 | -34.1368 | -11.3374
6.4095 | -34.1371 | -11.3476
6.4156 | -34.1374 | -11.3578
6.4217 | -34.1377 | -11.368
6.4278 | -34.1381 | -11.3783
6.4339 | -34.1384 | -11.3883
6.44 | -34.1387 | -11.3984
6.4461 | -34.139 | -11.4088
6.4522 | -34.1394 | -11.4188
6.4583 | -34.1396 | -11.4291
6.4644 | -34.1399 | -11.4394
6.4705 | -34.1402 | -11.4492
6.4766 | -34.1406 | -11.4594
6.4827 | -34.1409 | -11.4696
6.4888 | -34.1412 | -11.4796
6.4949 | -34.1415 | -11.4895
6.501 | -34.1418 | -11.4997
6.5071 | -34.1421 | -11.5096
6.5132 | -34.1425 | -11.5198
6.5194 | -34.1427 | -11.53
6.5255 | -34.1431 | -11.5399
6.5316 | -34.1434 | -11.5502
6.5377 | -34.1437 | -11.5603
6.5438 | -34.1439 | -11.5703
6.5499 | -34.1443 | -11.5803
6.556 | -34.1446 | -11.5904
6.5621 | -34.1449 | -11.6007
6.5682 | -34.1452 | -11.6109
6.5743 | -34.1456 | -11.6209
6.5804 | -34.1459 | -11.6311
6.5865 | -34.1462 | -11.6413
6.5926 | -34.1465 | -11.6513
6.5987 | -34.1468 | -11.6611
6.6048 | -34.1471 | -11.6713
6.6109 | -34.1475 | -11.6813
6.617 | -34.1477 | -11.6915
6.6231 | -34.148 | -11.7016
6.6292 | -34.1484 | -11.7117
6.6353 | -34.1487 | -11.7221
6.6414 | -34.149 | -11.7321
6.6475 | -34.1493 | -11.7421
6.6536 | -34.1496 | -11.7521
6.6597 | -34.1499 | -11.7619
6.6659 | -34.1503 | -11.772
6.672 | -34.1506 | -11.782
6.6781 | -34.1509 | -11.792
6.6842 | -34.1512 | -11.802
6.6903 | -34.1516 | -11.812
6.6964 | -34.1519 | -11.8221
6.7025 | -34.1522 | -11.832
6.7086 | -34.1525 | -11.8421
6.7147 | -34.1528 | -11.8522
6.7208 | -34.1531 | -11.8624
6.7269 | -34.1534 | -11.8724
6.733 | -34.1537 | -11.8822
6.7391 | -34.1541 | -11.8922
6.7452 | -34.1544 | -11.9021
6.7513 | -34.1547 | -11.9118
6.7574 | -34.155 | -11.922
6.7635 | -34.1553 | -11.9318
6.7696 | -34.1557 | -11.9417
6.7757 | -34.156 | -11.9517
6.7818 | -34.1563 | -11.9618
6.7879 | -34.1567 | -11.9716
6.794 | -34.157 | -11.9814
6.8001 | -34.1573 | -11.9914
6.8063 | -34.1576 | -12.0014
6.8124 | -34.1579 | -12.0115
6.8185 | -34.1582 | -12.0217
6.8246 | -34.1585 | -12.0314
6.8307 | -34.1588 | -12.0413
6.8368 | -34.1591 | -12.0515
6.8429 | -34.1595 | -12.0615
6.849 | -34.1598 | -12.0716
6.8551 | -34.1601 | -12.0814
6.8612 | -34.1605 | -12.0913
6.8673 | -34.1607 | -12.1013
6.8734 | -34.1611 | -12.1113
6.8795 | -34.1614 | -12.1213
6.8856 | -34.1617 | -12.1313
6.8917 | -34.162 | -12.1412
6.8978 | -34.1623 | -12.1509
6.9039 | -34.1626 | -12.1606
6.91 | -34.163 | -12.1707
6.9161 | -34.1633 | -12.1804
6.9222 | -34.1636 | -12.1903
6.9283 | -34.164 | -12.2
6.9344 | -34.1643 | -12.21
6.9405 | -34.1647 | -12.2199
6.9466 | -34.165 | -12.23
6.9528 | -34.1653 | -12.2401
6.9589 | -34.1657 | -12.2502
6.965 | -34.166 | -12.2602
6.9711 | -34.1664 | -12.2702
6.9772 | -34.1667 | -12.2802
6.9833 | -34.167 | -12.2901
6.9894 | -34.1674 | -12.3002
6.9955 | -34.1677 | -12.3099
7.0016 | -34.168 | -12.3198
7.0077 | -34.1684 | -12.3297
7.0138 | -34.1687 | -12.3396
7.0199 | -34.169 | -12.3494
7.026 | -34.1693 | -12.3593
7.0321 | -34.1697 | -12.369
7.0382 | -34.17 | -12.3786
7.0443 | -34.1704 | -12.3887
7.0504 | -34.1707 | -12.3987
7.0565 | -34.171 | -12.4086
7.0626 | -34.1714 | -12.4183
7.0687 | -34.1717 | -12.4278
7.0748 | -34.172 | -12.4374
7.0809 | -34.1724 | -12.4473
7.087 | -34.1727 | -12.4573
7.0932 | -34.173 | -12.4672
7.0993 | -34.1733 | -12.477
7.1054 | -34.1737 | -12.4868
7.1115 | -34.174 | -12.4964
7.1176 | -34.1744 | -12.5062
7.1237 | -34.1747 | -12.516
7.1298 | -34.175 | -12.5258
7.1359 | -34.1753 | -12.5357
7.142 | -34.1757 | -12.5453
7.1481 | -34.176 | -12.5551
7.1542 | -34.1763 | -12.565
7.1603 | -34.1767 | -12.5748
7.1664 | -34.177 | -12.5846
7.1725 | -34.1774 | -12.5945
7.1786 | -34.1778 | -12.6043
7.1847 | -34.1781 | -12.6142
7.1908 | -34.1784 | -12.6242
7.1969 | -34.1787 | -12.6342
7.203 | -34.179 | -12.6439
7.2091 | -34.1794 | -12.6536
7.2152 | -34.1797 | -12.6635
7.2213 | -34.1801 | -12.6731
7.2274 | -34.1804 | -12.6829
7.2335 | -34.1807 | -12.6927
7.2397 | -34.1811 | -12.7025
7.2458 | -34.1815 | -12.7121
7.2519 | -34.1818 | -12.7219
7.258 | -34.1821 | -12.7319
7.2641 | -34.1825 | -12.7417
7.2702 | -34.1828 | -12.7517
7.2763 | -34.1832 | -12.7615
7.2824 | -34.1835 | -12.7715
7.2885 | -34.1839 | -12.7813
7.2946 | -34.1842 | -12.7912
7.3007 | -34.1845 | -12.8011
7.3068 | -34.1848 | -12.8111
7.3129 | -34.1852 | -12.8212
7.319 | -34.1856 | -12.831
7.3251 | -34.1859 | -12.8406
7.3312 | -34.1862 | -12.8505
7.3373 | -34.1866 | -12.8601
7.3434 | -34.1869 | -12.8698
7.3495 | -34.1873 | -12.8795
7.3556 | -34.1877 | -12.8894
7.3617 | -34.188 | -12.8994
7.3678 | -34.1883 | -12.9092
7.3739 | -34.1887 | -12.9192
7.3801 | -34.189 | -12.9292
7.3862 | -34.1894 | -12.939
7.3923 | -34.1897 | -12.949
7.3984 | -34.1901 | -12.9586
7.4045 | -34.1905 | -12.9683
7.4106 | -34.1908 | -12.9781
7.4167 | -34.1912 | -12.988
7.4228 | -34.1916 | -12.998
7.4289 | -34.1919 | -13.0078
7.435 | -34.1922 | -13.0174
7.4411 | -34.1926 | -13.0274
7.4472 | -34.1929 | -13.0372
7.4533 | -34.1933 | -13.0469
7.4594 | -34.1937 | -13.057
7.4655 | -34.194 | -13.0666
7.4716 | -34.1944 | -13.0766
7.4777 | -34.1947 | -13.0864
7.4838 | -34.195 | -13.0961
7.4899 | -34.1954 | -13.1059
7.496 | -34.1957 | -13.1155
7.5021 | -34.1961 | -13.1253
7.5082 | -34.1964 | -13.1351
7.5143 | -34.1968 | -13.1449
7.5204 | -34.1971 | -13.1549
7.5266 | -34.1975 | -13.1646
7.5327 | -34.1978 | -13.1743
7.5388 | -34.1982 | -13.1844
7.5449 | -34.1986 | -13.1944
7.551 | -34.1989 | -13.2041
7.5571 | -34.1993 | -13.214
7.5632 | -34.1997 | -13.2239
7.5693 | -34.2 | -13.2336
7.5754 | -34.2004 | -13.2433
7.5815 | -34.2007 | -13.2531
7.5876 | -34.2011 | -13.2632
7.5937 | -34.2015 | -13.2732
7.5998 | -34.2018 | -13.2829
7.6059 | -34.2022 | -13.2929
7.612 | -34.2026 | -13.3027
7.6181 | -34.2029 | -13.3125
7.6242 | -34.2033 | -13.3226
7.6303 | -34.2037 | -13.3327
7.6364 | -34.204 | -13.3426
7.6425 | -34.2043 | -13.3524
7.6486 | -34.2047 | -13.3623
7.6547 | -34.205 | -13.372
7.6608 | -34.2054 | -13.3818
7.667 | -34.2057 | -13.3918
7.6731 | -34.2061 | -13.4017
7.6792 | -34.2065 | -13.4118
7.6853 | -34.2068 | -13.4217
7.6914 | -34.2072 | -13.4317
7.6975 | -34.2076 | -13.4416
7.7036 | -34.2079 | -13.4515
7.7097 | -34.2083 | -13.4616
7.7158 | -34.2087 | -13.4715
7.7219 | -34.2091 | -13.4813
7.728 | -34.2094 | -13.4913
7.7341 | -34.2098 | -13.5013
7.7402 | -34.2101 | -13.5112
7.7463 | -34.2105 | -13.5212
7.7524 | -34.2108 | -13.5313
7.7585 | -34.2112 | -13.5412
7.7646 | -34.2115 | -13.5511
7.7707 | -34.2119 | -13.5612
7.7768 | -34.2122 | -13.5712
7.7829 | -34.2126 | -13.5814
7.789 | -34.213 | -13.5916
7.7951 | -34.2133 | -13.6015
7.8012 | -34.2137 | -13.6113
7.8073 | -34.2141 | -13.6214
7.8135 | -34.2144 | -13.6313
7.8196 | -34.2148 | -13.6415
7.8257 | -34.2152 | -13.6516
7.8318 | -34.2155 | -13.6616
7.8379 | -34.2159 | -13.6715
7.844 | -34.2163 | -13.6813
7.8501 | -34.2166 | -13.6913
7.8562 | -34.217 | -13.7015
7.8623 | -34.2174 | -13.7113
7.8684 | -34.2177 | -13.7213
7.8745 | -34.2181 | -13.7313
7.8806 | -34.2184 | -13.7414
7.8867 | -34.2188 | -13.7511
7.8928 | -34.2192 | -13.7609
7.8989 | -34.2196 | -13.7708
7.905 | -34.2199 | -13.7809
7.9111 | -34.2203 | -13.7909
7.9172 | -34.2206 | -13.8008
7.9233 | -34.221 | -13.8109
7.9294 | -34.2214 | -13.8209
7.9355 | -34.2217 | -13.8307
7.9416 | -34.222 | -13.8407
7.9477 | -34.2224 | -13.8505
7.9539 | -34.2228 | -13.8604
7.96 | -34.2232 | -13.8705
7.9661 | -34.2235 | -13.8805
7.9722 | -34.2239 | -13.8903
7.9783 | -34.2243 | -13.9005
7.9844 | -34.2247 | -13.9104
7.9905 | -34.225 | -13.9205
7.9966 | -34.2254 | -13.9307
8.0027 | -34.2257 | -13.9406
8.0088 | -34.2261 | -13.9505
8.0149 | -34.2264 | -13.9605
8.021 | -34.2268 | -13.9706
8.0271 | -34.2272 | -13.9808
8.0332 | -34.2275 | -13.9908
8.0393 | -34.2279 | -14.0009
8.0454 | -34.2283 | -14.011
8.0515 | -34.2287 | -14.021
8.0576 | -34.2291 | -14.0313
8.0637 | -34.2294 | -14.0412
8.0698 | -34.2298 | -14.0511
8.0759 | -34.2301 | -14.061
8.082 | -34.2305 | -14.0708
8.0881 | -34.2308 | -14.0809
8.0942 | -34.2312 | -14.0909
8.1004 | -34.2316 | -14.101
8.1065 | -34.232 | -14.1108
8.1126 | -34.2323 | -14.1208
8.1187 | -34.2327 | -14.1309
8.1248 | -34.2331 | -14.1407
8.1309 | -34.2334 | -14.1505
8.137 | -34.2338 | -14.1603
8.1431 | -34.2342 | -14.1705
8.1492 | -34.2345 | -14.1805
8.1553 | -34.2349 | -14.1905
8.1614 | -34.2353 | -14.2006
8.1675 | -34.2356 | -14.2106
8.1736 | -34.236 | -14.2207
8.1797 | -34.2363 | -14.2308
8.1858 | -34.2367 | -14.2408
8.1919 | -34.2371 | -14.2508
8.198 | -34.2374 | -14.2606
8.2041 | -34.2378 | -14.2706
8.2102 | -34.2382 | -14.2806
8.2163 | -34.2385 | -14.2904
8.2224 | -34.2389 | -14.3004
8.2285 | -34.2393 | -14.3102
8.2346 | -34.2396 | -14.3201
8.2408 | -34.24 | -14.3299
8.2469 | -34.2404 | -14.34
8.253 | -34.2408 | -14.3498
8.2591 | -34.2411 | -14.3599
8.2652 | -34.2416 | -14.3699
8.2713 | -34.2419 | -14.3799
8.2774 | -34.2423 | -14.39
8.2835 | -34.2427 | -14.4002
8.2896 | -34.243 | -14.41
8.2957 | -34.2434 | -14.42
8.3018 | -34.2438 | -14.43
8.3079 | -34.2442 | -14.4398
8.314 | -34.2446 | -14.4498
8.3201 | -34.245 | -14.4596
8.3262 | -34.2453 | -14.4697
8.3323 | -34.2457 | -14.4796
8.3384 | -34.2461 | -14.4897
8.3445 | -34.2464 | -14.4996
8.3506 | -34.2468 | -14.5096
8.3567 | -34.2472 | -14.5197
8.3628 | -34.2475 | -14.5296
8.3689 | -34.2479 | -14.5395
8.375 | -34.2483 | -14.5495
8.3812 | -34.2486 | -14.5594
8.3873 | -34.249 | -14.5694
8.3934 | -34.2494 | -14.5794
8.3995 | -34.2497 | -14.5892
8.4056 | -34.2501 | -14.5991
8.4117 | -34.2505 | -14.6092
8.4178 | -34.2508 | -14.6192
8.4239 | -34.2512 | -14.629
8.43 | -34.2516 | -14.639
8.4361 | -34.2519 | -14.6491
8.4422 | -34.2523 | -14.6589
8.4483 | -34.2527 | -14.6687
8.4544 | -34.253 | -14.6787
8.4605 | -34.2534 | -14.6885
8.4666 | -34.2538 | -14.6987
8.4727 | -34.2541 | -14.7087
8.4788 | -34.2545 | -14.7186
8.4849 | -34.2549 | -14.7285
8.491 | -34.2552 | -14.7385
8.4971 | -34.2556 | -14.7482
8.5032 | -34.256 | -14.758
8.5093 | -34.2564 | -14.7677
8.5154 | -34.2567 | -14.7777
8.5215 | -34.2571 | -14.7876
8.5277 | -34.2575 | -14.7975
8.5338 | -34.2578 | -14.8074
8.5399 | -34.2582 | -14.8173
8.546 | -34.2586 | -14.8269
8.5521 | -34.259 | -14.8368
8.5582 | -34.2593 | -14.8467
8.5643 | -34.2597 | -14.8563
8.5704 | -34.2601 | -14.8659
8.5765 | -34.2605 | -14.8759
8.5826 | -34.2608 | -14.8855
8.5887 | -34.2612 | -14.8954
8.5948 | -34.2616 | -14.9053
8.6009 | -34.262 | -14.915
8.607 | -34.2624 | -14.925
8.6131 | -34.2627 | -14.9349
8.6192 | -34.2631 | -14.9445
8.6253 | -34.2635 | -14.9544
8.6314 | -34.2639 | -14.9641
8.6375 | -34.2642 | -14.974
8.6436 | -34.2646 | -14.9839
8.6497 | -34.265 | -14.9935
8.6558 | -34.2653 | -15.0033
8.6619 | -34.2657 | -15.0132
8.6681 | -34.266 | -15.0231
8.6742 | -34.2664 | -15.0329
8.6803 | -34.2668 | -15.0428
8.6864 | -34.2672 | -15.0523
8.6925 | -34.2675 | -15.062
8.6986 | -34.2679 | -15.0719
8.7047 | -34.2683 | -15.0818
8.7108 | -34.2686 | -15.0917
8.7169 | -34.269 | -15.1015
8.723 | -34.2694 | -15.1111
8.7291 | -34.2697 | -15.1207
8.7352 | -34.2701 | -15.1307
8.7413 | -34.2705 | -15.1406
8.7474 | -34.2708 | -15.1502
8.7535 | -34.2713 | -15.1601
8.7596 | -34.2716 | -15.1699
8.7657 | -34.272 | -15.1797
8.7718 | -34.2724 | -15.1896
8.7779 | -34.2728 | -15.1995
8.784 | -34.2731 | -15.2093
8.7901 | -34.2735 | -15.2191
8.7962 | -34.2739 | -15.2287
8.8023 | -34.2743 | -15.2384
8.8084 | -34.2746 | -15.2482
8.8146 | -34.275 | -15.2579
8.8207 | -34.2754 | -15.2677
8.8268 | -34.2758 | -15.2772
8.8329 | -34.2761 | -15.2869
8.839 | -34.2765 | -15.2967
8.8451 | -34.2769 | -15.3064
8.8512 | -34.2773 | -15.3161
8.8573 | -34.2776 | -15.3259
8.8634 | -34.278 | -15.3355
8.8695 | -34.2784 | -15.3454
8.8756 | -34.2787 | -15.3552
8.8817 | -34.2791 | -15.3648
8.8878 | -34.2795 | -15.3748
8.8939 | -34.2799 | -15.3845
8.9 | -34.2803 | -15.3943
8.9061 | -34.2806 | -15.404
8.9122 | -34.281 | -15.4137
8.9183 | -34.2814 | -15.4235
8.9244 | -34.2818 | -15.4334
8.9305 | -34.2822 | -15.4429
8.9366 | -34.2826 | -15.4526
8.9427 | -34.2829 | -15.4625
8.9488 | -34.2833 | -15.4721
8.955 | -34.2837 | -15.4818
8.9611 | -34.2841 | -15.4917
8.9672 | -34.2845 | -15.5012
8.9733 | -34.2849 | -15.5108
8.9794 | -34.2853 | -15.5204
8.9855 | -34.2857 | -15.5302
8.9916 | -34.286 | -15.5401
8.9977 | -34.2864 | -15.5498
9.0038 | -34.2868 | -15.5593
9.0099 | -34.2872 | -15.569
9.016 | -34.2876 | -15.5788
9.0221 | -34.288 | -15.5884
9.0282 | -34.2884 | -15.598
9.0343 | -34.2888 | -15.6077
9.0404 | -34.2892 | -15.6174
9.0465 | -34.2896 | -15.6273
9.0526 | -34.29 | -15.6369
9.0587 | -34.2904 | -15.6467
9.0648 | -34.2908 | -15.6563
9.0709 | -34.2912 | -15.6659
9.077 | -34.2916 | -15.6756
9.0831 | -34.292 | -15.6854
9.0892 | -34.2924 | -15.6952
9.0953 | -34.2928 | -15.7049
9.1015 | -34.2932 | -15.7146
9.1076 | -34.2936 | -15.7241
9.1137 | -34.294 | -15.7338
9.1198 | -34.2944 | -15.7434
9.1259 | -34.2947 | -15.7531
9.132 | -34.2951 | -15.7629
9.1381 | -34.2956 | -15.7724
9.1442 | -34.296 | -15.7818
9.1503 | -34.2964 | -15.7915
9.1564 | -34.2968 | -15.8011
9.1625 | -34.2971 | -15.8108
9.1686 | -34.2976 | -15.8204
9.1747 | -34.298 | -15.83
9.1808 | -34.2984 | -15.8396
9.1869 | -34.2988 | -15.8492
9.193 | -34.2992 | -15.8588
9.1991 | -34.2996 | -15.8685
9.2052 | -34.3 | -15.8781
9.2113 | -34.3004 | -15.8878
9.2174 | -34.3008 | -15.8973
9.2235 | -34.3012 | -15.9069
9.2296 | -34.3017 | -15.9165
9.2357 | -34.3021 | -15.9263
9.2419 | -34.3025 | -15.9359
9.248 | -34.3029 | -15.9454
9.2541 | -34.3033 | -15.9551
9.2602 | -34.3037 | -15.9649
9.2663 | -34.3041 | -15.9743
9.2724 | -34.3045 | -15.9838
9.2785 | -34.3049 | -15.9933
9.2846 | -34.3054 | -16.0029
9.2907 | -34.3058 | -16.0125
9.2968 | -34.3062 | -16.0222
9.3029 | -34.3066 | -16.0319
9.309 | -34.307 | -16.0416
9.3151 | -34.3074 | -16.0511
9.3212 | -34.3078 | -16.0605
9.3273 | -34.3082 | -16.07
9.3334 | -34.3086 | -16.0795
9.3395 | -34.309 | -16.089
9.3456 | -34.3094 | -16.0984
9.3517 | -34.3098 | -16.108
9.3578 | -34.3102 | -16.1176
9.3639 | -34.3107 | -16.127
9.37 | -34.3111 | -16.1366
9.3761 | -34.3115 | -16.1463
9.3822 | -34.3119 | -16.1559
9.3884 | -34.3123 | -16.1652
9.3945 | -34.3128 | -16.1748
9.4006 | -34.3132 | -16.1844
9.4067 | -34.3136 | -16.1941
9.4128 | -34.314 | -16.2038
9.4189 | -34.3144 | -16.2135
9.425 | -34.3149 | -16.2229
9.4311 | -34.3153 | -16.2325
9.4372 | -34.3157 | -16.242
9.4433 | -34.3161 | -16.2516
9.4494 | -34.3166 | -16.261
9.4555 | -34.317 | -16.2706
9.4616 | -34.3174 | -16.2801
9.4677 | -34.3178 | -16.2897
9.4738 | -34.3183 | -16.2993
9.4799 | -34.3187 | -16.3089
9.486 | -34.3191 | -16.3184
9.4921 | -34.3195 | -16.328
9.4982 | -34.32 | -16.3373
9.5043 | -34.3204 | -16.3468
9.5104 | -34.3208 | -16.3561
9.5165 | -34.3212 | -16.3656
9.5226 | -34.3216 | -16.3753
9.5288 | -34.3221 | -16.3846
9.5349 | -34.3225 | -16.3942
9.541 | -34.3229 | -16.4035
9.5471 | -34.3233 | -16.4129
9.5532 | -34.3237 | -16.4225
9.5593 | -34.3241 | -16.4322
9.5654 | -34.3246 | -16.4416
9.5715 | -34.325 | -16.4511
9.5776 | -34.3254 | -16.4606
9.5837 | -34.3258 | -16.4702
9.5898 | -34.3261 | -16.4799
9.5959 | -34.3266 | -16.4893
9.602 | -34.327 | -16.4987
9.6081 | -34.3274 | -16.5081
9.6142 | -34.3279 | -16.5175
9.6203 | -34.3283 | -16.5271
9.6264 | -34.3287 | -16.5365
9.6325 | -34.3291 | -16.5461
9.6386 | -34.3295 | -16.5553
9.6447 | -34.3299 | -16.5649
9.6508 | -34.3303 | -16.5743
9.6569 | -34.3307 | -16.5837
9.663 | -34.3312 | -16.5933
9.6691 | -34.3316 | -16.603
9.6753 | -34.332 | -16.6125
9.6814 | -34.3325 | -16.6218
9.6875 | -34.3329 | -16.6311
9.6936 | -34.3333 | -16.6405
9.6997 | -34.3337 | -16.6502
9.7058 | -34.3342 | -16.6596
9.7119 | -34.3346 | -16.6693
9.718 | -34.335 | -16.6788
9.7241 | -34.3355 | -16.6882
9.7302 | -34.3359 | -16.6976
9.7363 | -34.3363 | -16.7069
9.7424 | -34.3367 | -16.7164
9.7485 | -34.3371 | -16.7258
9.7546 | -34.3375 | -16.735
9.7607 | -34.338 | -16.7447
9.7668 | -34.3384 | -16.7542
9.7729 | -34.3388 | -16.7638
9.779 | -34.3393 | -16.7729
9.7851 | -34.3397 | -16.7824
9.7912 | -34.3402 | -16.7916
9.7973 | -34.3406 | -16.8013
9.8034 | -34.341 | -16.8105
9.8095 | -34.3414 | -16.8204
9.8157 | -34.3419 | -16.8298
9.8218 | -34.3423 | -16.8393
9.8279 | -34.3427 | -16.8484
9.834 | -34.343 | -16.8582
9.8401 | -34.3435 | -16.8676
9.8462 | -34.3439 | -16.8768
9.8523 | -34.3443 | -16.8862
9.8584 | -34.3448 | -16.8956
9.8645 | -34.3452 | -16.905
9.8706 | -34.3456 | -16.9146
9.8767 | -34.346 | -16.9239
9.8828 | -34.3464 | -16.933
9.8889 | -34.3468 | -16.9423
9.895 | -34.3472 | -16.9517
9.9011 | -34.3476 | -16.9612
9.9072 | -34.3481 | -16.9706
9.9133 | -34.3485 | -16.98
9.9194 | -34.3489 | -16.9895
9.9255 | -34.3494 | -16.999
9.9316 | -34.3498 | -17.0084
9.9377 | -34.3503 | -17.0179
9.9438 | -34.3507 | -17.0273
9.9499 | -34.3511 | -17.037
9.956 | -34.3515 | -17.0463
9.9622 | -34.352 | -17.0556
9.9683 | -34.3524 | -17.0651
9.9744 | -34.3528 | -17.0747
9.9805 | -34.3532 | -17.0841
9.9866 | -34.3536 | -17.0933
9.9927 | -34.354 | -17.1026
9.9988 | -34.3544 | -17.1123
10.0049 | -34.3549 | -17.1217
10.011 | -34.3553 | -17.1313
10.0171 | -34.3557 | -17.1406
10.0232 | -34.3562 | -17.1498
10.0293 | -34.3566 | -17.1594
10.0354 | -34.3571 | -17.1687
10.0415 | -34.3575 | -17.1782
10.0476 | -34.3579 | -17.1878
10.0537 | -34.3583 | -17.1972
10.0598 | -34.3587 | -17.2068
10.0659 | -34.3591 | -17.2162
10.072 | -34.3595 | -17.2253
10.0781 | -34.36 | -17.2346
10.0842 | -34.3604 | -17.244
10.0903 | -34.3608 | -17.2535
10.0964 | -34.3612 | -17.2627
10.1026 | -34.3616 | -17.2722
10.1087 | -34.362 | -17.2816
10.1148 | -34.3624 | -17.2911
10.1209 | -34.3628 | -17.3004
10.127 | -34.3633 | -17.3097
10.1331 | -34.3637 | -17.3191
10.1392 | -34.3641 | -17.3285
10.1453 | -34.3646 | -17.3379
10.1514 | -34.365 | -17.3471
10.1575 | -34.3654 | -17.3564
10.1636 | -34.3658 | -17.3655
10.1697 | -34.3662 | -17.3748
10.1758 | -34.3666 | -17.384
10.1819 | -34.367 | -17.3934
10.188 | -34.3674 | -17.4029
10.1941 | -34.3678 | -17.4123
10.2002 | -34.3682 | -17.4214
10.2063 | -34.3687 | -17.4306
10.2124 | -34.3691 | -17.4399
10.2185 | -34.3695 | -17.4494
10.2246 | -34.3699 | -17.4586
10.2307 | -34.3703 | -17.4679
10.2368 | -34.3707 | -17.4768
10.2429 | -34.3711 | -17.4861
10.2491 | -34.3715 | -17.4957
10.2552 | -34.372 | -17.5051
10.2613 | -34.3724 | -17.5145
10.2674 | -34.3728 | -17.5236
10.2735 | -34.3732 | -17.5332
10.2796 | -34.3737 | -17.5423
10.2857 | -34.3741 | -17.5518
10.2918 | -34.3745 | -17.5613
10.2979 | -34.375 | -17.5705
10.304 | -34.3754 | -17.5797
10.3101 | -34.3758 | -17.5892
10.3162 | -34.3762 | -17.5986
10.3223 | -34.3766 | -17.6077
10.3284 | -34.377 | -17.6171
10.3345 | -34.3774 | -17.6264
10.3406 | -34.3778 | -17.6357
10.3467 | -34.3782 | -17.645
10.3528 | -34.3786 | -17.6543
10.3589 | -34.379 | -17.6636
10.365 | -34.3794 | -17.6731
10.3711 | -34.3799 | -17.6825
10.3772 | -34.3803 | -17.6915
10.3833 | -34.3807 | -17.7009
10.3895 | -34.3811 | -17.7101
10.3956 | -34.3816 | -17.7195
10.4017 | -34.382 | -17.7289
10.4078 | -34.3824 | -17.7379
10.4139 | -34.3828 | -17.7475
10.42 | -34.3832 | -17.7569
10.4261 | -34.3836 | -17.7661
10.4322 | -34.384 | -17.7753
10.4383 | -34.3845 | -17.7846
10.4444 | -34.3849 | -17.7936
10.4505 | -34.3853 | -17.8031
10.4566 | -34.3857 | -17.8123
10.4627 | -34.3861 | -17.8215
10.4688 | -34.3865 | -17.8307
10.4749 | -34.3869 | -17.8402
10.481 | -34.3873 | -17.8497
10.4871 | -34.3877 | -17.8589
10.4932 | -34.3882 | -17.8681
10.4993 | -34.3886 | -17.8774
10.5054 | -34.389 | -17.8867
10.5115 | -34.3894 | -17.8959
10.5176 | -34.3898 | -17.9053
10.5237 | -34.3902 | -17.9146
10.5298 | -34.3907 | -17.9238
10.536 | -34.391 | -17.9331
10.5421 | -34.3915 | -17.9424
10.5482 | -34.3919 | -17.9517
10.5543 | -34.3923 | -17.9606
10.5604 | -34.3927 | -17.9698
10.5665 | -34.3931 | -17.9789
10.5726 | -34.3935 | -17.9883
10.5787 | -34.3939 | -17.9974
10.5848 | -34.3943 | -18.0064
10.5909 | -34.3947 | -18.0157
10.597 | -34.3952 | -18.0251
10.6031 | -34.3956 | -18.0345
10.6092 | -34.396 | -18.0438
10.6153 | -34.3964 | -18.0532
10.6214 | -34.3968 | -18.0624
10.6275 | -34.3972 | -18.0715
10.6336 | -34.3976 | -18.0809
10.6397 | -34.398 | -18.0903
10.6458 | -34.3984 | -18.0996
10.6519 | -34.3988 | -18.109
10.658 | -34.3992 | -18.1184
10.6641 | -34.3996 | -18.1276
10.6702 | -34.4 | -18.1369
10.6764 | -34.4005 | -18.1463
10.6825 | -34.4009 | -18.1554
10.6886 | -34.4013 | -18.1646
10.6947 | -34.4017 | -18.1738
10.7008 | -34.4021 | -18.1831
10.7069 | -34.4025 | -18.1924
10.713 | -34.4029 | -18.2018
10.7191 | -34.4033 | -18.211
10.7252 | -34.4037 | -18.22
10.7313 | -34.4041 | -18.2291
10.7374 | -34.4045 | -18.2383
10.7435 | -34.4049 | -18.2475
10.7496 | -34.4053 | -18.2566
10.7557 | -34.4057 | -18.2658
10.7618 | -34.4062 | -18.2747
10.7679 | -34.4066 | -18.2839
10.774 | -34.407 | -18.2933
10.7801 | -34.4074 | -18.3025
10.7862 | -34.4078 | -18.3115
10.7923 | -34.4082 | -18.3209
10.7984 | -34.4087 | -18.3301
10.8045 | -34.4091 | -18.3392
10.8106 | -34.4095 | -18.3487
10.8168 | -34.4099 | -18.3579
10.8229 | -34.4103 | -18.3672
10.829 | -34.4107 | -18.3763
10.8351 | -34.4111 | -18.3854
10.8412 | -34.4116 | -18.3945
10.8473 | -34.412 | -18.4036
10.8534 | -34.4124 | -18.4128
10.8595 | -34.4129 | -18.422
10.8656 | -34.4133 | -18.4309
10.8717 | -34.4137 | -18.4402
10.8778 | -34.4141 | -18.4494
10.8839 | -34.4145 | -18.4586
10.89 | -34.4149 | -18.4678
10.8961 | -34.4154 | -18.477
10.9022 | -34.4158 | -18.4861
10.9083 | -34.4162 | -18.4951
10.9144 | -34.4166 | -18.5042
10.9205 | -34.417 | -18.5132
10.9266 | -34.4174 | -18.5224
10.9327 | -34.4178 | -18.5315
10.9388 | -34.4183 | -18.5405
10.9449 | -34.4187 | -18.5496
10.951 | -34.4191 | -18.5588
10.9571 | -34.4195 | -18.568
10.9633 | -34.42 | -18.5772
10.9694 | -34.4204 | -18.5865
10.9755 | -34.4208 | -18.5957
10.9816 | -34.4212 | -18.6048
10.9877 | -34.4216 | -18.6139
10.9938 | -34.422 | -18.6232
10.9999 | -34.4225 | -18.6321
11.006 | -34.4229 | -18.6411
11.0121 | -34.4233 | -18.6502
11.0182 | -34.4238 | -18.6594
11.0243 | -34.4242 | -18.6686
11.0304 | -34.4246 | -18.6779
11.0365 | -34.425 | -18.6872
11.0426 | -34.4254 | -18.6963
11.0487 | -34.4258 | -18.7057
11.0548 | -34.4263 | -18.7146
11.0609 | -34.4267 | -18.7238
11.067 | -34.4271 | -18.7329
11.0731 | -34.4276 | -18.7419
11.0792 | -34.428 | -18.7509
11.0853 | -34.4284 | -18.7601
11.0914 | -34.4289 | -18.7695
11.0975 | -34.4293 | -18.7786
11.1037 | -34.4297 | -18.7877
11.1098 | -34.4301 | -18.7969
11.1159 | -34.4306 | -18.8058
11.122 | -34.431 | -18.8149
11.1281 | -34.4314 | -18.824
11.1342 | -34.4318 | -18.833
11.1403 | -34.4323 | -18.8423
11.1464 | -34.4327 | -18.8515
11.1525 | -34.4331 | -18.8606
11.1586 | -34.4335 | -18.8699
11.1647 | -34.434 | -18.8789
11.1708 | -34.4344 | -18.888
11.1769 | -34.4348 | -18.8973
11.183 | -34.4352 | -18.9062
11.1891 | -34.4357 | -18.9153
11.1952 | -34.4361 | -18.9245
11.2013 | -34.4365 | -18.9337
11.2074 | -34.437 | -18.9428
11.2135 | -34.4374 | -18.9521
11.2196 | -34.4378 | -18.9613
11.2257 | -34.4383 | -18.9705
11.2318 | -34.4387 | -18.9796
11.2379 | -34.4392 | -18.9889
11.244 | -34.4396 | -18.9982
11.2502 | -34.4401 | -19.0072
11.2563 | -34.4405 | -19.0164
11.2624 | -34.4409 | -19.0255
11.2685 | -34.4414 | -19.0347
11.2746 | -34.4418 | -19.0437
11.2807 | -34.4422 | -19.053
11.2868 | -34.4426 | -19.0623
11.2929 | -34.4431 | -19.071
11.299 | -34.4436 | -19.0802
11.3051 | -34.444 | -19.0895
11.3112 | -34.4444 | -19.0987
11.3173 | -34.4448 | -19.1076
11.3234 | -34.4453 | -19.1164
11.3295 | -34.4457 | -19.1255
11.3356 | -34.4462 | -19.1346
11.3417 | -34.4466 | -19.1435
11.3478 | -34.447 | -19.1529
11.3539 | -34.4474 | -19.1622
11.36 | -34.4478 | -19.171
11.3661 | -34.4483 | -19.1799
11.3722 | -34.4487 | -19.1889
11.3783 | -34.4492 | -19.1981
11.3844 | -34.4496 | -19.2072
11.3906 | -34.4501 | -19.2163
11.3967 | -34.4505 | -19.2253
11.4028 | -34.451 | -19.2345
11.4089 | -34.4514 | -19.2436
11.415 | -34.4518 | -19.2527
11.4211 | -34.4522 | -19.2618
11.4272 | -34.4527 | -19.2709
11.4333 | -34.4531 | -19.2798
11.4394 | -34.4535 | -19.2889
11.4455 | -34.4539 | -19.2979
11.4516 | -34.4544 | -19.3071
11.4577 | -34.4548 | -19.3161
11.4638 | -34.4553 | -19.325
11.4699 | -34.4557 | -19.3341
11.476 | -34.4561 | -19.3434
11.4821 | -34.4566 | -19.3525
11.4882 | -34.4571 | -19.3618
11.4943 | -34.4575 | -19.371
11.5004 | -34.458 | -19.3799
11.5065 | -34.4584 | -19.3888
11.5126 | -34.4589 | -19.3977
11.5187 | -34.4593 | -19.4068
11.5248 | -34.4597 | -19.416
11.5309 | -34.4602 | -19.4251
11.5371 | -34.4607 | -19.4339
11.5432 | -34.4611 | -19.4431
11.5493 | -34.4615 | -19.4523
11.5554 | -34.462 | -19.4614
11.5615 | -34.4624 | -19.4706
11.5676 | -34.4629 | -19.4798
11.5737 | -34.4634 | -19.4887
11.5798 | -34.4638 | -19.4976
11.5859 | -34.4643 | -19.5066
11.592 | -34.4647 | -19.5155
11.5981 | -34.4652 | -19.5246
11.6042 | -34.4656 | -19.5335
11.6103 | -34.4661 | -19.5425
11.6164 | -34.4665 | -19.5515
11.6225 | -34.4669 | -19.5605
11.6286 | -34.4674 | -19.5696
11.6347 | -34.4678 | -19.5788
11.6408 | -34.4683 | -19.5875
11.6469 | -34.4687 | -19.5965
11.653 | -34.4692 | -19.6054
11.6591 | -34.4696 | -19.6144
11.6652 | -34.4701 | -19.6234
11.6713 | -34.4705 | -19.6324
11.6775 | -34.471 | -19.6412
11.6836 | -34.4715 | -19.6502
11.6897 | -34.472 | -19.6592
11.6958 | -34.4724 | -19.6684
11.7019 | -34.4728 | -19.6775
11.708 | -34.4733 | -19.6865
11.7141 | -34.4738 | -19.6954
11.7202 | -34.4743 | -19.7045
11.7263 | -34.4747 | -19.7135
11.7324 | -34.4751 | -19.7227
11.7385 | -34.4756 | -19.7318
11.7446 | -34.476 | -19.7407
11.7507 | -34.4765 | -19.7496
11.7568 | -34.477 | -19.7587
11.7629 | -34.4774 | -19.7678
11.769 | -34.4778 | -19.7766
11.7751 | -34.4783 | -19.7856
11.7812 | -34.4788 | -19.7945
11.7873 | -34.4793 | -19.8036
11.7934 | -34.4797 | -19.8126
11.7995 | -34.4802 | -19.8215
11.8056 | -34.4806 | -19.8305
11.8117 | -34.4811 | -19.8395
11.8178 | -34.4816 | -19.8483
11.824 | -34.482 | -19.8571
11.8301 | -34.4825 | -19.866
11.8362 | -34.483 | -19.8748
11.8423 | -34.4834 | -19.8839
11.8484 | -34.4839 | -19.8928
11.8545 | -34.4844 | -19.9018
11.8606 | -34.4848 | -19.9106
11.8667 | -34.4853 | -19.9197
11.8728 | -34.4857 | -19.9284
11.8789 | -34.4862 | -19.9374
11.885 | -34.4866 | -19.9464
11.8911 | -34.4871 | -19.9553
11.8972 | -34.4876 | -19.9642
11.9033 | -34.488 | -19.9731
11.9094 | -34.4885 | -19.982
11.9155 | -34.489 | -19.9909
11.9216 | -34.4895 | -19.9998
11.9277 | -34.4899 | -20.0087
11.9338 | -34.4904 | -20.0176
11.9399 | -34.4908 | -20.0267
11.946 | -34.4913 | -20.0356
11.9521 | -34.4918 | -20.0446
11.9582 | -34.4923 | -20.0535
11.9644 | -34.4928 | -20.0623
11.9705 | -34.4932 | -20.0712
11.9766 | -34.4937 | -20.0801
11.9827 | -34.4942 | -20.089
11.9888 | -34.4947 | -20.0979
11.9949 | -34.4952 | -20.1065
12.001 | -34.4957 | -20.1155
12.0071 | -34.4961 | -20.1245
12.0132 | -34.4966 | -20.1332
12.0193 | -34.497 | -20.1422
12.0254 | -34.4975 | -20.1512
12.0315 | -34.498 | -20.16
12.0376 | -34.4984 | -20.1688
12.0437 | -34.4989 | -20.1777
12.0498 | -34.4994 | -20.1867
12.0559 | -34.4998 | -20.1955
12.062 | -34.5003 | -20.2041
12.0681 | -34.5008 | -20.213
12.0742 | -34.5013 | -20.2219
12.0803 | -34.5017 | -20.2309
12.0864 | -34.5022 | -20.2398
12.0925 | -34.5027 | -20.2487
12.0986 | -34.5032 | -20.2575
12.1047 | -34.5036 | -20.2664
12.1109 | -34.5041 | -20.2751
12.117 | -34.5046 | -20.284
12.1231 | -34.505 | -20.2929
12.1292 | -34.5055 | -20.3017
12.1353 | -34.506 | -20.3105
12.1414 | -34.5065 | -20.3194
12.1475 | -34.5069 | -20.3285
12.1536 | -34.5074 | -20.3375
12.1597 | -34.5078 | -20.3463
12.1658 | -34.5083 | -20.3551
12.1719 | -34.5088 | -20.3638
12.178 | -34.5093 | -20.3727
12.1841 | -34.5098 | -20.3816
12.1902 | -34.5102 | -20.3905
12.1963 | -34.5107 | -20.3994
12.2024 | -34.5112 | -20.4082
12.2085 | -34.5117 | -20.417
12.2146 | -34.5122 | -20.426
12.2207 | -34.5126 | -20.4348
12.2268 | -34.5131 | -20.4436
12.2329 | -34.5136 | -20.4522
12.239 | -34.5141 | -20.4609
12.2451 | -34.5146 | -20.4698
12.2513 | -34.515 | -20.4788
12.2574 | -34.5155 | -20.4877
12.2635 | -34.516 | -20.4964
12.2696 | -34.5165 | -20.5053
12.2757 | -34.517 | -20.5139
12.2818 | -34.5175 | -20.5228
12.2879 | -34.5179 | -20.5316
12.294 | -34.5184 | -20.5404
12.3001 | -34.5189 | -20.5492
12.3062 | -34.5194 | -20.558
12.3123 | -34.5199 | -20.5666
12.3184 | -34.5204 | -20.5753
12.3245 | -34.5209 | -20.584
12.3306 | -34.5213 | -20.5928
12.3367 | -34.5218 | -20.6016
12.3428 | -34.5223 | -20.6105
12.3489 | -34.5228 | -20.6189
12.355 | -34.5232 | -20.6279
12.3611 | -34.5237 | -20.6367
12.3672 | -34.5242 | -20.6454
12.3733 | -34.5247 | -20.6542
12.3794 | -34.5252 | -20.6627
12.3855 | -34.5256 | -20.6713
12.3916 | -34.5261 | -20.6802
12.3978 | -34.5266 | -20.6891
12.4039 | -34.5271 | -20.6979
12.41 | -34.5276 | -20.7067
12.4161 | -34.5281 | -20.7151
12.4222 | -34.5286 | -20.7238
12.4283 | -34.529 | -20.7324
12.4344 | -34.5295 | -20.7412
12.4405 | -34.53 | -20.7501
12.4466 | -34.5304 | -20.7587
12.4527 | -34.5309 | -20.7674
12.4588 | -34.5314 | -20.7762
12.4649 | -34.5319 | -20.7849
12.471 | -34.5324 | -20.7939
12.4771 | -34.5328 | -20.8026
12.4832 | -34.5333 | -20.8114
12.4893 | -34.5338 | -20.8199
12.4954 | -34.5343 | -20.8284
12.5015 | -34.5348 | -20.837
12.5076 | -34.5353 | -20.8459
12.5137 | -34.5358 | -20.8547
12.5198 | -34.5362 | -20.8633
12.5259 | -34.5367 | -20.8721
12.532 | -34.5372 | -20.8808
12.5382 | -34.5377 | -20.8895
12.5443 | -34.5381 | -20.8982
12.5504 | -34.5386 | -20.9071
12.5565 | -34.5391 | -20.9157
12.5626 | -34.5396 | -20.9246
12.5687 | -34.5401 | -20.9334
12.5748 | -34.5405 | -20.9422
12.5809 | -34.541 | -20.9509
12.587 | -34.5415 | -20.9597
12.5931 | -34.542 | -20.9683
12.5992 | -34.5425 | -20.9768
12.6053 | -34.543 | -20.9855
12.6114 | -34.5435 | -20.9941
12.6175 | -34.544 | -21.003
12.6236 | -34.5444 | -21.0118
12.6297 | -34.545 | -21.0204
12.6358 | -34.5454 | -21.0289
12.6419 | -34.5459 | -21.0376
12.648 | -34.5464 | -21.0465
12.6541 | -34.5469 | -21.0551
12.6602 | -34.5474 | -21.0635
12.6663 | -34.5478 | -21.0722
12.6724 | -34.5483 | -21.0808
12.6785 | -34.5488 | -21.0895
12.6847 | -34.5493 | -21.0981
12.6908 | -34.5498 | -21.1068
12.6969 | -34.5503 | -21.1154
12.703 | -34.5508 | -21.1241
12.7091 | -34.5512 | -21.1329
12.7152 | -34.5517 | -21.1414
12.7213 | -34.5522 | -21.1502
12.7274 | -34.5527 | -21.159
12.7335 | -34.5531 | -21.1674
12.7396 | -34.5536 | -21.176
12.7457 | -34.5541 | -21.1848
12.7518 | -34.5545 | -21.1935
12.7579 | -34.555 | -21.2022
12.764 | -34.5555 | -21.2109
12.7701 | -34.556 | -21.2196
12.7762 | -34.5565 | -21.2282
12.7823 | -34.557 | -21.2368
12.7884 | -34.5575 | -21.2454
12.7945 | -34.5579 | -21.2542
12.8006 | -34.5584 | -21.2628
12.8067 | -34.5589 | -21.2714
12.8128 | -34.5593 | -21.2799
12.8189 | -34.5598 | -21.2886
12.8251 | -34.5603 | -21.2972
12.8312 | -34.5608 | -21.3061
12.8373 | -34.5612 | -21.3146
12.8434 | -34.5617 | -21.3231
12.8495 | -34.5622 | -21.3317
12.8556 | -34.5627 | -21.3401
12.8617 | -34.5631 | -21.3488
12.8678 | -34.5636 | -21.3576
12.8739 | -34.5641 | -21.3662
12.88 | -34.5645 | -21.3747
12.8861 | -34.565 | -21.3833
12.8922 | -34.5654 | -21.3921
12.8983 | -34.5659 | -21.4007
12.9044 | -34.5664 | -21.4092
12.9105 | -34.5669 | -21.4178
12.9166 | -34.5673 | -21.4263
12.9227 | -34.5678 | -21.4349
12.9288 | -34.5683 | -21.4435
12.9349 | -34.5688 | -21.4521
12.941 | -34.5693 | -21.4607
12.9471 | -34.5697 | -21.4693
12.9532 | -34.5702 | -21.4779
12.9593 | -34.5707 | -21.4867
12.9654 | -34.5712 | -21.4954
12.9716 | -34.5717 | -21.5042
12.9777 | -34.5722 | -21.5124
12.9838 | -34.5726 | -21.5212
12.9899 | -34.5731 | -21.5297
12.996 | -34.5735 | -21.5382
13.0021 | -34.574 | -21.5468
13.0082 | -34.5745 | -21.5555
13.0143 | -34.575 | -21.5637
13.0204 | -34.5754 | -21.5723
13.0265 | -34.5759 | -21.5809
13.0326 | -34.5764 | -21.5893
13.0387 | -34.5768 | -21.5981
13.0448 | -34.5773 | -21.6067
13.0509 | -34.5777 | -21.6151
13.057 | -34.5782 | -21.6236
13.0631 | -34.5787 | -21.6319
13.0692 | -34.5792 | -21.6403
13.0753 | -34.5797 | -21.6489
13.0814 | -34.5801 | -21.6575
13.0875 | -34.5806 | -21.6659
13.0936 | -34.581 | -21.6745
13.0997 | -34.5815 | -21.6828
13.1058 | -34.5819 | -21.6913
13.112 | -34.5824 | -21.7
13.1181 | -34.5829 | -21.7085
13.1242 | -34.5833 | -21.7169
13.1303 | -34.5838 | -21.7255
13.1364 | -34.5844 | -21.7344
13.1425 | -34.5848 | -21.7428
13.1486 | -34.5853 | -21.7511
13.1547 | -34.5857 | -21.7596
13.1608 | -34.5862 | -21.7683
13.1669 | -34.5866 | -21.7768
13.173 | -34.5871 | -21.7856
13.1791 | -34.5876 | -21.7941
13.1852 | -34.588 | -21.8027
13.1913 | -34.5885 | -21.8111
13.1974 | -34.589 | -21.8198
13.2035 | -34.5895 | -21.8283
13.2096 | -34.59 | -21.8369
13.2157 | -34.5905 | -21.8452
13.2218 | -34.5909 | -21.8537
13.2279 | -34.5914 | -21.8625
13.234 | -34.5919 | -21.8708
13.2401 | -34.5923 | -21.8793
13.2462 | -34.5928 | -21.8878
13.2524 | -34.5933 | -21.8963
13.2585 | -34.5937 | -21.9045
13.2646 | -34.5942 | -21.9131
13.2707 | -34.5946 | -21.9217
13.2768 | -34.5951 | -21.9303
13.2829 | -34.5955 | -21.9389
13.289 | -34.596 | -21.9472
13.2951 | -34.5965 | -21.9554
13.3012 | -34.597 | -21.9638
13.3073 | -34.5975 | -21.9725
13.3134 | -34.5979 | -21.981
13.3195 | -34.5984 | -21.9895
13.3256 | -34.5988 | -21.9979
13.3317 | -34.5993 | -22.0063
13.3378 | -34.5997 | -22.015
13.3439 | -34.6002 | -22.0235
13.35 | -34.6007 | -22.0322
13.3561 | -34.6012 | -22.0405
13.3622 | -34.6016 | -22.0489
13.3683 | -34.602 | -22.0575
13.3744 | -34.6025 | -22.0659
13.3805 | -34.603 | -22.0742
13.3866 | -34.6034 | -22.0826
13.3927 | -34.6038 | -22.0911
13.3989 | -34.6043 | -22.0996
13.405 | -34.6047 | -22.108
13.4111 | -34.6052 | -22.1163
13.4172 | -34.6056 | -22.1246
13.4233 | -34.6061 | -22.1332
13.4294 | -34.6065 | -22.1417
13.4355 | -34.607 | -22.15
13.4416 | -34.6074 | -22.158
13.4477 | -34.6079 | -22.1666
13.4538 | -34.6083 | -22.1751
13.4599 | -34.6088 | -22.1835
13.466 | -34.6092 | -22.192
13.4721 | -34.6097 | -22.2007
13.4782 | -34.6101 | -22.2092
13.4843 | -34.6106 | -22.2177
13.4904 | -34.611 | -22.2261
13.4965 | -34.6115 | -22.2347
13.5026 | -34.6119 | -22.2433
13.5087 | -34.6124 | -22.2513
13.5148 | -34.6128 | -22.26
13.5209 | -34.6133 | -22.2687
13.527 | -34.6137 | -22.2772
13.5331 | -34.6142 | -22.2855
13.5393 | -34.6147 | -22.2938
13.5454 | -34.6151 | -22.3023
13.5515 | -34.6156 | -22.3107
13.5576 | -34.6161 | -22.3191
13.5637 | -34.6165 | -22.3277
13.5698 | -34.617 | -22.3363
13.5759 | -34.6175 | -22.3443
13.582 | -34.6179 | -22.3528
13.5881 | -34.6184 | -22.3611
13.5942 | -34.6189 | -22.3698
13.6003 | -34.6193 | -22.3781
13.6064 | -34.6198 | -22.3865
13.6125 | -34.6202 | -22.3949
13.6186 | -34.6207 | -22.4034
13.6247 | -34.6212 | -22.4119
13.6308 | -34.6217 | -22.4201
13.6369 | -34.6221 | -22.4286
13.643 | -34.6225 | -22.437
13.6491 | -34.6229 | -22.4454
13.6552 | -34.6234 | -22.4539
13.6613 | -34.6239 | -22.4623
13.6674 | -34.6243 | -22.4707
13.6735 | -34.6247 | -22.4792
13.6796 | -34.6252 | -22.4875
13.6858 | -34.6256 | -22.4957
13.6919 | -34.6261 | -22.5042
13.698 | -34.6266 | -22.5126
13.7041 | -34.627 | -22.5212
13.7102 | -34.6275 | -22.5296
13.7163 | -34.6279 | -22.5379
13.7224 | -34.6284 | -22.5462
13.7285 | -34.6288 | -22.5546
13.7346 | -34.6293 | -22.563
13.7407 | -34.6298 | -22.5711
13.7468 | -34.6302 | -22.5796
13.7529 | -34.6306 | -22.5878
13.759 | -34.6311 | -22.5962
13.7651 | -34.6315 | -22.6046
13.7712 | -34.632 | -22.6132
13.7773 | -34.6324 | -22.6215
13.7834 | -34.6329 | -22.63
13.7895 | -34.6333 | -22.6382
13.7956 | -34.6338 | -22.6465
13.8017 | -34.6342 | -22.6548
13.8078 | -34.6347 | -22.6632
13.8139 | -34.6351 | -22.6717
13.82 | -34.6355 | -22.68
13.8262 | -34.636 | -22.6883
13.8323 | -34.6365 | -22.697
13.8384 | -34.6369 | -22.7054
13.8445 | -34.6374 | -22.7137
13.8506 | -34.6378 | -22.7221
13.8567 | -34.6383 | -22.7305
13.8628 | -34.6387 | -22.7391
13.8689 | -34.6392 | -22.7475
13.875 | -34.6396 | -22.7557
13.8811 | -34.64 | -22.7641
13.8872 | -34.6405 | -22.7724
13.8933 | -34.6409 | -22.7805
13.8994 | -34.6413 | -22.789
13.9055 | -34.6418 | -22.7976
13.9116 | -34.6423 | -22.806
13.9177 | -34.6427 | -22.8144
13.9238 | -34.6432 | -22.8227
13.9299 | -34.6436 | -22.8309
13.936 | -34.644 | -22.8393
13.9421 | -34.6445 | -22.8473
13.9482 | -34.6449 | -22.8557
13.9543 | -34.6454 | -22.8643
13.9604 | -34.6458 | -22.8728
13.9665 | -34.6463 | -22.8811
13.9727 | -34.6468 | -22.8897
13.9788 | -34.6473 | -22.8979
13.9849 | -34.6477 | -22.9063
13.991 | -34.6482 | -22.9148
13.9971 | -34.6486 | -22.9231
14.0032 | -34.649 | -22.9315
14.0093 | -34.6495 | -22.9399
14.0154 | -34.6499 | -22.9481
14.0215 | -34.6503 | -22.9565
14.0276 | -34.6508 | -22.9648
14.0337 | -34.6513 | -22.9729
14.0398 | -34.6518 | -22.9812
14.0459 | -34.6522 | -22.9894
14.052 | -34.6526 | -22.998
14.0581 | -34.6531 | -23.0063
14.0642 | -34.6535 | -23.0147
14.0703 | -34.654 | -23.0231
14.0764 | -34.6544 | -23.0312
14.0825 | -34.6549 | -23.0395
14.0886 | -34.6553 | -23.0479
14.0947 | -34.6558 | -23.0562
14.1008 | -34.6562 | -23.0645
14.1069 | -34.6567 | -23.0727
14.1131 | -34.6571 | -23.0809
14.1192 | -34.6575 | -23.0891
14.1253 | -34.6579 | -23.0972
14.1314 | -34.6584 | -23.1055
14.1375 | -34.6588 | -23.1137
14.1436 | -34.6593 | -23.1219
14.1497 | -34.6598 | -23.1302
14.1558 | -34.6602 | -23.1386
14.1619 | -34.6606 | -23.1469
14.168 | -34.6611 | -23.1551
14.1741 | -34.6616 | -23.1632
14.1802 | -34.662 | -23.1716
14.1863 | -34.6625 | -23.1798
14.1924 | -34.6629 | -23.1881
14.1985 | -34.6634 | -23.1965
14.2046 | -34.6639 | -23.2049
14.2107 | -34.6643 | -23.213
14.2168 | -34.6648 | -23.2212
14.2229 | -34.6652 | -23.2294
14.229 | -34.6657 | -23.2377
14.2351 | -34.6661 | -23.246
14.2412 | -34.6665 | -23.2545
14.2473 | -34.667 | -23.2627
14.2534 | -34.6674 | -23.2708
14.2596 | -34.6679 | -23.2789
14.2657 | -34.6683 | -23.2873
14.2718 | -34.6687 | -23.2955
14.2779 | -34.6692 | -23.3037
14.284 | -34.6697 | -23.312
14.2901 | -34.6701 | -23.3205
14.2962 | -34.6706 | -23.3286
14.3023 | -34.671 | -23.3368
14.3084 | -34.6715 | -23.3451
14.3145 | -34.6719 | -23.3533
14.3206 | -34.6724 | -23.3618
14.3267 | -34.6729 | -23.3701
14.3328 | -34.6733 | -23.3784
14.3389 | -34.6738 | -23.3865
14.345 | -34.6742 | -23.3948
14.3511 | -34.6747 | -23.4031
14.3572 | -34.6751 | -23.4114
14.3633 | -34.6756 | -23.4197
14.3694 | -34.676 | -23.428
14.3755 | -34.6765 | -23.4362
14.3816 | -34.677 | -23.4446
14.3877 | -34.6775 | -23.4526
14.3938 | -34.6779 | -23.4611
14.4 | -34.6783 | -23.4693
14.4061 | -34.6788 | -23.4776
14.4122 | -34.6792 | -23.4857
14.4183 | -34.6797 | -23.4939
14.4244 | -34.6801 | -23.5021
14.4305 | -34.6806 | -23.5104
14.4366 | -34.6811 | -23.5189
14.4427 | -34.6816 | -23.5272
14.4488 | -34.682 | -23.5356
14.4549 | -34.6825 | -23.5438
14.461 | -34.6829 | -23.5521
14.4671 | -34.6834 | -23.5604
14.4732 | -34.6838 | -23.5688
14.4793 | -34.6843 | -23.5769
14.4854 | -34.6848 | -23.5852
14.4915 | -34.6852 | -23.5931
14.4976 | -34.6857 | -23.6012
14.5037 | -34.6861 | -23.6096
14.5098 | -34.6866 | -23.6181
14.5159 | -34.687 | -23.6263
14.522 | -34.6874 | -23.6346
14.5281 | -34.6879 | -23.6425
14.5342 | -34.6884 | -23.6509
14.5403 | -34.6889 | -23.659
14.5465 | -34.6893 | -23.6674
14.5526 | -34.6898 | -23.6759
14.5587 | -34.6902 | -23.6844
14.5648 | -34.6907 | -23.6924
14.5709 | -34.6912 | -23.7005
14.577 | -34.6916 | -23.7086
14.5831 | -34.692 | -23.7169
14.5892 | -34.6925 | -23.7251
14.5953 | -34.6929 | -23.7333
14.6014 | -34.6934 | -23.7417
14.6075 | -34.6938 | -23.7499
14.6136 | -34.6943 | -23.7583
14.6197 | -34.6947 | -23.7665
14.6258 | -34.6952 | -23.7748
14.6319 | -34.6956 | -23.7829
14.638 | -34.696 | -23.7912
14.6441 | -34.6965 | -23.7995
14.6502 | -34.6969 | -23.8078
14.6563 | -34.6974 | -23.816
14.6624 | -34.6979 | -23.8243
14.6685 | -34.6984 | -23.8323
14.6746 | -34.6988 | -23.8406
14.6807 | -34.6993 | -23.8487
14.6869 | -34.6998 | -23.8569
14.693 | -34.7002 | -23.865
14.6991 | -34.7007 | -23.8732
14.7052 | -34.7012 | -23.8812
14.7113 | -34.7017 | -23.8895
14.7174 | -34.7021 | -23.8978
14.7235 | -34.7026 | -23.9059
14.7296 | -34.703 | -23.914
14.7357 | -34.7035 | -23.922
14.7418 | -34.704 | -23.9301
14.7479 | -34.7045 | -23.9383
14.754 | -34.7049 | -23.9466
14.7601 | -34.7054 | -23.9546
14.7662 | -34.7059 | -23.9629
14.7723 | -34.7063 | -23.9708
14.7784 | -34.7068 | -23.9789
14.7845 | -34.7072 | -23.9869
14.7906 | -34.7077 | -23.9953
14.7967 | -34.7082 | -24.0036
14.8028 | -34.7086 | -24.0116
14.8089 | -34.7091 | -24.0197
14.815 | -34.7096 | -24.0278
14.8211 | -34.7101 | -24.036
14.8272 | -34.7105 | -24.0443
14.8334 | -34.711 | -24.0525
14.8395 | -34.7115 | -24.0605
14.8456 | -34.7119 | -24.0686
14.8517 | -34.7124 | -24.077
14.8578 | -34.7129 | -24.0851
14.8639 | -34.7133 | -24.0933
14.87 | -34.7138 | -24.1014
14.8761 | -34.7143 | -24.1096
14.8822 | -34.7148 | -24.1177
14.8883 | -34.7152 | -24.1258
14.8944 | -34.7157 | -24.1337
14.9005 | -34.7161 | -24.1417
14.9066 | -34.7166 | -24.15
14.9127 | -34.7171 | -24.158
14.9188 | -34.7176 | -24.1661
14.9249 | -34.718 | -24.1744
14.931 | -34.7185 | -24.1826
14.9371 | -34.7189 | -24.1908
14.9432 | -34.7194 | -24.1988
14.9493 | -34.7199 | -24.2068
14.9554 | -34.7204 | -24.2148
14.9615 | -34.7208 | -24.223
14.9676 | -34.7213 | -24.2313
14.9738 | -34.7218 | -24.2394
14.9799 | -34.7222 | -24.2476
14.986 | -34.7227 | -24.2557
14.9921 | -34.7232 | -24.2638
14.9982 | -34.7236 | -24.2721
15.0043 | -34.7241 | -24.2802
15.0104 | -34.7246 | -24.2885
15.0165 | -34.7251 | -24.2968
15.0226 | -34.7256 | -24.3047
15.0287 | -34.7261 | -24.3129
15.0348 | -34.7266 | -24.3211
15.0409 | -34.727 | -24.329
15.047 | -34.7275 | -24.3373
15.0531 | -34.728 | -24.3452
15.0592 | -34.7285 | -24.3534
15.0653 | -34.729 | -24.3616
15.0714 | -34.7294 | -24.3696
15.0775 | -34.7299 | -24.3776
15.0836 | -34.7304 | -24.3857
15.0897 | -34.7309 | -24.3937
15.0958 | -34.7314 | -24.4017
15.1019 | -34.7319 | -24.4098
15.108 | -34.7324 | -24.4179
15.1141 | -34.7328 | -24.4261
15.1203 | -34.7333 | -24.4342
15.1264 | -34.7338 | -24.4422
15.1325 | -34.7343 | -24.4502
15.1386 | -34.7348 | -24.4584
15.1447 | -34.7353 | -24.4665
15.1508 | -34.7358 | -24.4746
15.1569 | -34.7363 | -24.4829
15.163 | -34.7368 | -24.4908
15.1691 | -34.7373 | -24.4989
15.1752 | -34.7377 | -24.507
15.1813 | -34.7382 | -24.515
15.1874 | -34.7387 | -24.5231
15.1935 | -34.7392 | -24.531
15.1996 | -34.7397 | -24.5391
15.2057 | -34.7402 | -24.5472
15.2118 | -34.7406 | -24.5554
15.2179 | -34.7411 | -24.5635
15.224 | -34.7416 | -24.5715
15.2301 | -34.7421 | -24.5794
15.2362 | -34.7426 | -24.5876
15.2423 | -34.7431 | -24.5958
15.2484 | -34.7436 | -24.604
15.2545 | -34.7441 | -24.6122
15.2607 | -34.7446 | -24.6203
15.2668 | -34.7451 | -24.6281
15.2729 | -34.7456 | -24.6361
15.279 | -34.7461 | -24.6444
15.2851 | -34.7466 | -24.6524
15.2912 | -34.747 | -24.6605
15.2973 | -34.7475 | -24.6688
15.3034 | -34.748 | -24.6768
15.3095 | -34.7486 | -24.6849
15.3156 | -34.749 | -24.6928
15.3217 | -34.7496 | -24.7009
15.3278 | -34.75 | -24.7092
15.3339 | -34.7506 | -24.717
15.34 | -34.7511 | -24.7249
15.3461 | -34.7515 | -24.7329
15.3522 | -34.752 | -24.7409
15.3583 | -34.7525 | -24.749
15.3644 | -34.753 | -24.7568
15.3705 | -34.7535 | -24.7645
15.3766 | -34.754 | -24.7726
15.3827 | -34.7545 | -24.7805
15.3888 | -34.755 | -24.7886
15.3949 | -34.7555 | -24.7966
15.401 | -34.756 | -24.8047
15.4072 | -34.7565 | -24.8126
15.4133 | -34.757 | -24.8206
15.4194 | -34.7575 | -24.8286
15.4255 | -34.758 | -24.8366
15.4316 | -34.7584 | -24.8446
15.4377 | -34.759 | -24.8529
15.4438 | -34.7595 | -24.8607
15.4499 | -34.76 | -24.8687
15.456 | -34.7605 | -24.8766
15.4621 | -34.761 | -24.8846
15.4682 | -34.7615 | -24.8926
15.4743 | -34.7621 | -24.9004
15.4804 | -34.7625 | -24.9085
15.4865 | -34.763 | -24.9165
15.4926 | -34.7635 | -24.9245
15.4987 | -34.764 | -24.9326
15.5048 | -34.7645 | -24.9406
15.5109 | -34.7651 | -24.9486
15.517 | -34.7656 | -24.9566
15.5231 | -34.7661 | -24.9647
15.5292 | -34.7666 | -24.9727
15.5353 | -34.7671 | -24.9807
15.5414 | -34.7676 | -24.9886
15.5476 | -34.7681 | -24.9964
15.5537 | -34.7686 | -25.0045
15.5598 | -34.7691 | -25.0126
15.5659 | -34.7696 | -25.0206
15.572 | -34.7702 | -25.0286
15.5781 | -34.7707 | -25.0365
15.5842 | -34.7712 | -25.0444
15.5903 | -34.7717 | -25.0523
15.5964 | -34.7722 | -25.0602
15.6025 | -34.7728 | -25.0681
15.6086 | -34.7733 | -25.0762
15.6147 | -34.7738 | -25.0842
15.6208 | -34.7743 | -25.092
15.6269 | -34.7748 | -25.0999
15.633 | -34.7753 | -25.108
15.6391 | -34.7759 | -25.1159
15.6452 | -34.7763 | -25.1239
15.6513 | -34.7769 | -25.1316
15.6574 | -34.7774 | -25.1397
15.6635 | -34.7779 | -25.1476
15.6696 | -34.7785 | -25.1555
15.6757 | -34.779 | -25.1635
15.6818 | -34.7795 | -25.1712
15.688 | -34.78 | -25.1793
15.6941 | -34.7805 | -25.1872
15.7002 | -34.781 | -25.1952
15.7063 | -34.7815 | -25.2029
15.7124 | -34.782 | -25.2111
15.7185 | -34.7825 | -25.219
15.7246 | -34.783 | -25.2268
15.7307 | -34.7836 | -25.2348
15.7368 | -34.7841 | -25.2428
15.7429 | -34.7846 | -25.2507
15.749 | -34.7851 | -25.2587
15.7551 | -34.7856 | -25.2667
15.7612 | -34.7862 | -25.2745
15.7673 | -34.7867 | -25.2823
15.7734 | -34.7872 | -25.2903
15.7795 | -34.7877 | -25.2982
15.7856 | -34.7882 | -25.3061
15.7917 | -34.7888 | -25.3139
15.7978 | -34.7893 | -25.322
15.8039 | -34.7898 | -25.3299
15.81 | -34.7903 | -25.3379
15.8161 | -34.7908 | -25.3459
15.8222 | -34.7913 | -25.3539
15.8283 | -34.7919 | -25.3616
15.8345 | -34.7924 | -25.3694
15.8406 | -34.7929 | -25.3773
15.8467 | -34.7935 | -25.3851
15.8528 | -34.794 | -25.3931
15.8589 | -34.7945 | -25.401
15.865 | -34.7951 | -25.4087
15.8711 | -34.7956 | -25.4167
15.8772 | -34.7961 | -25.4245
15.8833 | -34.7966 | -25.4325
15.8894 | -34.7972 | -25.4405
15.8955 | -34.7977 | -25.4483
15.9016 | -34.7982 | -25.4561
15.9077 | -34.7987 | -25.464
15.9138 | -34.7992 | -25.472
15.9199 | -34.7997 | -25.4797
15.926 | -34.8003 | -25.4876
15.9321 | -34.8008 | -25.4954
15.9382 | -34.8013 | -25.5031
15.9443 | -34.8018 | -25.5109
15.9504 | -34.8024 | -25.5188
15.9565 | -34.8029 | -25.5266
15.9626 | -34.8034 | -25.5345
15.9687 | -34.8039 | -25.5421
15.9749 | -34.8044 | -25.5499
15.981 | -34.8049 | -25.5577
15.9871 | -34.8055 | -25.5657
15.9932 | -34.806 | -25.5735
15.9993 | -34.8065 | -25.5813
16.0054 | -34.8071 | -25.5888
16.0115 | -34.8076 | -25.5966
16.0176 | -34.8082 | -25.6045
16.0237 | -34.8087 | -25.6124
16.0298 | -34.8092 | -25.6201
16.0359 | -34.8097 | -25.6281
16.042 | -34.8103 | -25.6358
16.0481 | -34.8108 | -25.6437
16.0542 | -34.8113 | -25.6516
16.0603 | -34.8118 | -25.6594
16.0664 | -34.8124 | -25.6672
16.0725 | -34.8129 | -25.6748
16.0786 | -34.8134 | -25.6825
16.0847 | -34.8139 | -25.6905
16.0908 | -34.8145 | -25.6984
16.0969 | -34.815 | -25.7063
16.103 | -34.8155 | -25.7142
16.1091 | -34.816 | -25.7219
16.1152 | -34.8165 | -25.7298
16.1214 | -34.817 | -25.7377
16.1275 | -34.8176 | -25.7454
16.1336 | -34.8182 | -25.7532
16.1397 | -34.8187 | -25.761
16.1458 | -34.8192 | -25.7687
16.1519 | -34.8198 | -25.7767
16.158 | -34.8203 | -25.7846
16.1641 | -34.8208 | -25.7924
16.1702 | -34.8213 | -25.8002
16.1763 | -34.8219 | -25.8079
16.1824 | -34.8224 | -25.8155
16.1885 | -34.8229 | -25.8232
16.1946 | -34.8235 | -25.831
16.2007 | -34.824 | -25.8388
16.2068 | -34.8246 | -25.8465
16.2129 | -34.8251 | -25.8544
16.219 | -34.8257 | -25.862
16.2251 | -34.8262 | -25.8697
16.2312 | -34.8267 | -25.8775
16.2373 | -34.8272 | -25.8852
16.2434 | -34.8278 | -25.893
16.2495 | -34.8283 | -25.9006
16.2556 | -34.8289 | -25.9085
16.2618 | -34.8294 | -25.9162
16.2679 | -34.83 | -25.9241
16.274 | -34.8305 | -25.9317
16.2801 | -34.831 | -25.9396
16.2862 | -34.8315 | -25.9474
16.2923 | -34.8321 | -25.9552
16.2984 | -34.8326 | -25.963
16.3045 | -34.8331 | -25.9708
16.3106 | -34.8337 | -25.9785
16.3167 | -34.8342 | -25.9862
16.3228 | -34.8347 | -25.9938
16.3289 | -34.8352 | -26.0014
16.335 | -34.8358 | -26.0091
16.3411 | -34.8363 | -26.0169
16.3472 | -34.8368 | -26.0248
16.3533 | -34.8374 | -26.0326
16.3594 | -34.8379 | -26.0402
16.3655 | -34.8384 | -26.048
16.3716 | -34.839 | -26.0557
16.3777 | -34.8395 | -26.0635
16.3838 | -34.84 | -26.0711
16.3899 | -34.8406 | -26.0788
16.396 | -34.8411 | -26.0865
16.4021 | -34.8416 | -26.0942
16.4083 | -34.8421 | -26.1019
16.4144 | -34.8427 | -26.1098
16.4205 | -34.8432 | -26.1174
16.4266 | -34.8438 | -26.1249
16.4327 | -34.8443 | -26.1326
16.4388 | -34.8448 | -26.1403
16.4449 | -34.8453 | -26.1481
16.451 | -34.8459 | -26.1559
16.4571 | -34.8464 | -26.1638
16.4632 | -34.8469 | -26.1712
16.4693 | -34.8474 | -26.179
16.4754 | -34.848 | -26.1868
16.4815 | -34.8485 | -26.1945
16.4876 | -34.849 | -26.2022
16.4937 | -34.8495 | -26.2099
16.4998 | -34.85 | -26.2174
16.5059 | -34.8506 | -26.225
16.512 | -34.8511 | -26.2328
16.5181 | -34.8517 | -26.2407
16.5242 | -34.8522 | -26.2484
16.5303 | -34.8527 | -26.2559
16.5364 | -34.8532 | -26.2634
16.5425 | -34.8537 | -26.2713
16.5487 | -34.8542 | -26.2789
16.5548 | -34.8548 | -26.2867
16.5609 | -34.8553 | -26.2945
16.567 | -34.8558 | -26.3021
16.5731 | -34.8564 | -26.31
16.5792 | -34.8569 | -26.3176
16.5853 | -34.8574 | -26.3253
16.5914 | -34.8579 | -26.3331
16.5975 | -34.8584 | -26.3409
16.6036 | -34.859 | -26.3486
16.6097 | -34.8595 | -26.3565
16.6158 | -34.86 | -26.3641
16.6219 | -34.8605 | -26.3717
16.628 | -34.861 | -26.3795
16.6341 | -34.8615 | -26.3871
16.6402 | -34.8621 | -26.3946
16.6463 | -34.8626 | -26.4022
16.6524 | -34.8631 | -26.4098
16.6585 | -34.8637 | -26.4174
16.6646 | -34.8642 | -26.425
16.6707 | -34.8647 | -26.4325
16.6768 | -34.8653 | -26.4402
16.6829 | -34.8658 | -26.4478
16.689 | -34.8663 | -26.4553
16.6952 | -34.8668 | -26.4629
16.7013 | -34.8673 | -26.4705
16.7074 | -34.8679 | -26.4777
16.7135 | -34.8684 | -26.4852
16.7196 | -34.8689 | -26.4927
16.7257 | -34.8695 | -26.5005
16.7318 | -34.87 | -26.5081
16.7379 | -34.8705 | -26.5156
16.744 | -34.8711 | -26.5231
16.7501 | -34.8716 | -26.5308
16.7562 | -34.8721 | -26.5384
16.7623 | -34.8726 | -26.5458
16.7684 | -34.8731 | -26.5536
16.7745 | -34.8737 | -26.5612
16.7806 | -34.8742 | -26.5685
16.7867 | -34.8747 | -26.5761
16.7928 | -34.8753 | -26.5838
16.7989 | -34.8758 | -26.5914
16.805 | -34.8763 | -26.5991
16.8111 | -34.8769 | -26.6066
16.8172 | -34.8774 | -26.6143
16.8233 | -34.8779 | -26.6221
16.8294 | -34.8784 | -26.6297
16.8356 | -34.8789 | -26.6373
16.8417 | -34.8795 | -26.6449
16.8478 | -34.88 | -26.6522
16.8539 | -34.8805 | -26.6597
16.86 | -34.881 | -26.6674
16.8661 | -34.8815 | -26.6749
16.8722 | -34.8821 | -26.6827
16.8783 | -34.8826 | -26.6903
16.8844 | -34.8831 | -26.6977
16.8905 | -34.8837 | -26.7052
16.8966 | -34.8842 | -26.7127
16.9027 | -34.8847 | -26.7202
16.9088 | -34.8852 | -26.7279
16.9149 | -34.8857 | -26.7356
16.921 | -34.8863 | -26.7431
16.9271 | -34.8868 | -26.7507
16.9332 | -34.8873 | -26.7583
16.9393 | -34.8878 | -26.7659
16.9454 | -34.8883 | -26.7734
16.9515 | -34.8888 | -26.7811
16.9576 | -34.8893 | -26.7884
16.9637 | -34.8898 | -26.7959
16.9698 | -34.8903 | -26.8034
16.9759 | -34.8909 | -26.8109
16.9821 | -34.8914 | -26.8185
16.9882 | -34.8919 | -26.8258
16.9943 | -34.8924 | -26.8332
17.0004 | -34.8929 | -26.8408
17.0065 | -34.8934 | -26.8483
17.0126 | -34.894 | -26.8559
17.0187 | -34.8945 | -26.8635
17.0248 | -34.895 | -26.8709
17.0309 | -34.8955 | -26.8785
17.037 | -34.896 | -26.886
17.0431 | -34.8966 | -26.8936
17.0492 | -34.8971 | -26.9011
17.0553 | -34.8976 | -26.9086
17.0614 | -34.8981 | -26.9161
17.0675 | -34.8986 | -26.9237
17.0736 | -34.8991 | -26.9313
17.0797 | -34.8997 | -26.9388
17.0858 | -34.9002 | -26.9463
17.0919 | -34.9007 | -26.954
17.098 | -34.9012 | -26.9611
17.1041 | -34.9017 | -26.9686
17.1102 | -34.9023 | -26.9762
17.1163 | -34.9028 | -26.9837
17.1225 | -34.9033 | -26.9912
17.1286 | -34.9039 | -26.9984
17.1347 | -34.9044 | -27.0059
17.1408 | -34.9049 | -27.0136
17.1469 | -34.9054 | -27.0211
17.153 | -34.9059 | -27.0286
17.1591 | -34.9064 | -27.0361
17.1652 | -34.907 | -27.0432
17.1713 | -34.9075 | -27.0507
17.1774 | -34.908 | -27.0582
17.1835 | -34.9085 | -27.0658
17.1896 | -34.909 | -27.0732
17.1957 | -34.9096 | -27.0807
17.2018 | -34.9101 | -27.088
17.2079 | -34.9106 | -27.0955
17.214 | -34.9112 | -27.1031
17.2201 | -34.9117 | -27.1106
17.2262 | -34.9122 | -27.1183
17.2323 | -34.9127 | -27.1259
17.2384 | -34.9132 | -27.1334
17.2445 | -34.9137 | -27.1408
17.2506 | -34.9142 | -27.1483
17.2567 | -34.9147 | -27.1558
17.2628 | -34.9152 | -27.1632
17.269 | -34.9158 | -27.1705
17.2751 | -34.9163 | -27.178
17.2812 | -34.9168 | -27.1856
17.2873 | -34.9173 | -27.1933
17.2934 | -34.9178 | -27.2007
17.2995 | -34.9183 | -27.2083
17.3056 | -34.9188 | -27.2155
17.3117 | -34.9193 | -27.2231
17.3178 | -34.9198 | -27.2306
17.3239 | -34.9203 | -27.238
17.33 | -34.9208 | -27.2455
17.3361 | -34.9213 | -27.2529
17.3422 | -34.9218 | -27.2602
17.3483 | -34.9223 | -27.2675
17.3544 | -34.9229 | -27.2749
17.3605 | -34.9234 | -27.2823
17.3666 | -34.9239 | -27.2896
17.3727 | -34.9244 | -27.297
17.3788 | -34.9249 | -27.3044
17.3849 | -34.9254 | -27.3119
17.391 | -34.9259 | -27.3194
17.3971 | -34.9264 | -27.3267
17.4032 | -34.9269 | -27.3341
17.4094 | -34.9274 | -27.3413
17.4155 | -34.9279 | -27.3486
17.4216 | -34.9284 | -27.356
17.4277 | -34.9289 | -27.3635
17.4338 | -34.9294 | -27.371
17.4399 | -34.9299 | -27.3785
17.446 | -34.9304 | -27.3856
17.4521 | -34.9309 | -27.3931
17.4582 | -34.9314 | -27.4004
17.4643 | -34.9319 | -27.4078
17.4704 | -34.9324 | -27.4152
17.4765 | -34.9329 | -27.4227
17.4826 | -34.9334 | -27.4301
17.4887 | -34.9339 | -27.4376
17.4948 | -34.9344 | -27.4451
17.5009 | -34.9349 | -27.4525
17.507 | -34.9354 | -27.46
17.5131 | -34.9359 | -27.4674
17.5192 | -34.9364 | -27.4746
17.5253 | -34.9369 | -27.482
17.5314 | -34.9373 | -27.4895
17.5375 | -34.9379 | -27.4969
17.5436 | -34.9384 | -27.5044
17.5497 | -34.9389 | -27.5119
17.5559 | -34.9394 | -27.5192
17.562 | -34.9399 | -27.5266
17.5681 | -34.9404 | -27.534
17.5742 | -34.9409 | -27.5416
17.5803 | -34.9414 | -27.5488
17.5864 | -34.9419 | -27.5562
17.5925 | -34.9424 | -27.5635
17.5986 | -34.943 | -27.5709
17.6047 | -34.9435 | -27.5784
17.6108 | -34.9439 | -27.5857
17.6169 | -34.9444 | -27.593
17.623 | -34.9449 | -27.6002
17.6291 | -34.9454 | -27.6074
17.6352 | -34.9459 | -27.6147
17.6413 | -34.9464 | -27.6222
17.6474 | -34.9469 | -27.6296
17.6535 | -34.9474 | -27.6369
17.6596 | -34.9479 | -27.6441
17.6657 | -34.9484 | -27.6515
17.6718 | -34.9489 | -27.659
17.6779 | -34.9494 | -27.6662
17.684 | -34.9499 | -27.6735
17.6901 | -34.9504 | -27.6809
17.6963 | -34.9509 | -27.6881
17.7024 | -34.9514 | -27.6954
17.7085 | -34.9519 | -27.7027
17.7146 | -34.9524 | -27.71
17.7207 | -34.9529 | -27.7172
17.7268 | -34.9534 | -27.7244
17.7329 | -34.9539 | -27.7318
17.739 | -34.9544 | -27.7391
17.7451 | -34.9549 | -27.7464
17.7512 | -34.9554 | -27.7538
17.7573 | -34.9559 | -27.7614
17.7634 | -34.9564 | -27.7685
17.7695 | -34.9569 | -27.7759
17.7756 | -34.9574 | -27.7832
17.7817 | -34.9578 | -27.7906
17.7878 | -34.9584 | -27.7981
17.7939 | -34.9588 | -27.8053
17.8 | -34.9594 | -27.8125
17.8061 | -34.9599 | -27.8199
17.8122 | -34.9604 | -27.8272
17.8183 | -34.9608 | -27.8346
17.8244 | -34.9614 | -27.8419
17.8305 | -34.9618 | -27.849
17.8366 | -34.9623 | -27.8561
17.8428 | -34.9628 | -27.8635
17.8489 | -34.9633 | -27.8708
17.855 | -34.9637 | -27.878
17.8611 | -34.9642 | -27.8853
17.8672 | -34.9647 | -27.8925
17.8733 | -34.9652 | -27.8999
17.8794 | -34.9657 | -27.9072
17.8855 | -34.9662 | -27.9144
17.8916 | -34.9666 | -27.9217
17.8977 | -34.9671 | -27.9289
17.9038 | -34.9676 | -27.9359
17.9099 | -34.9681 | -27.9432
17.916 | -34.9686 | -27.9505
17.9221 | -34.969 | -27.9579
17.9282 | -34.9695 | -27.965
17.9343 | -34.97 | -27.9722
17.9404 | -34.9705 | -27.9793
17.9465 | -34.971 | -27.9866
17.9526 | -34.9715 | -27.9939
17.9587 | -34.9719 | -28.0013
17.9648 | -34.9724 | -28.0085
17.9709 | -34.9729 | -28.0158
17.977 | -34.9734 | -28.0229
17.9832 | -34.9739 | -28.0302
17.9893 | -34.9744 | -28.0376
17.9954 | -34.9749 | -28.0448
18.0015 | -34.9753 | -28.0521
18.0076 | -34.9758 | -28.0592
18.0137 | -34.9763 | -28.0666
18.0198 | -34.9768 | -28.0737
18.0259 | -34.9773 | -28.0808
18.032 | -34.9778 | -28.088
18.0381 | -34.9783 | -28.0952
18.0442 | -34.9787 | -28.1023
18.0503 | -34.9792 | -28.1097
18.0564 | -34.9797 | -28.117
18.0625 | -34.9802 | -28.1243
18.0686 | -34.9807 | -28.1316
18.0747 | -34.9811 | -28.1389
18.0808 | -34.9816 | -28.146
18.0869 | -34.9821 | -28.1533
18.093 | -34.9826 | -28.1605
18.0991 | -34.9831 | -28.1678
18.1052 | -34.9835 | -28.1751
18.1113 | -34.984 | -28.1823
18.1174 | -34.9845 | -28.1895
18.1236 | -34.985 | -28.1967
18.1297 | -34.9855 | -28.204
18.1358 | -34.9859 | -28.2112
18.1419 | -34.9864 | -28.2185
18.148 | -34.9869 | -28.2256
18.1541 | -34.9874 | -28.2325
18.1602 | -34.9879 | -28.2398
18.1663 | -34.9884 | -28.247
18.1724 | -34.9888 | -28.2543
18.1785 | -34.9893 | -28.2616
18.1846 | -34.9898 | -28.2687
18.1907 | -34.9903 | -28.2759
18.1968 | -34.9908 | -28.2831
18.2029 | -34.9913 | -28.2904
18.209 | -34.9918 | -28.2978
18.2151 | -34.9922 | -28.3051
18.2212 | -34.9927 | -28.3122
18.2273 | -34.9932 | -28.3196
18.2334 | -34.9937 | -28.3267
18.2395 | -34.9942 | -28.3339
18.2456 | -34.9947 | -28.3412
18.2517 | -34.9951 | -28.3485
18.2578 | -34.9956 | -28.3556
18.2639 | -34.9961 | -28.3629
18.2701 | -34.9966 | -28.3703
18.2762 | -34.9971 | -28.3773
18.2823 | -34.9976 | -28.3845
18.2884 | -34.9981 | -28.3915
18.2945 | -34.9986 | -28.3984
18.3006 | -34.9991 | -28.4057
18.3067 | -34.9995 | -28.413
18.3128 | -35.0001 | -28.4202
18.3189 | -35.0005 | -28.4273
18.325 | -35.001 | -28.4343
18.3311 | -35.0014 | -28.4416
18.3372 | -35.0019 | -28.4488
18.3433 | -35.0024 | -28.4561
18.3494 | -35.0028 | -28.4633
18.3555 | -35.0033 | -28.4706
18.3616 | -35.0038 | -28.4774
18.3677 | -35.0043 | -28.4846
18.3738 | -35.0048 | -28.4918
18.3799 | -35.0052 | -28.499
18.386 | -35.0057 | -28.5062
18.3921 | -35.0062 | -28.5133
18.3982 | -35.0067 | -28.5202
18.4043 | -35.0071 | -28.5273
18.4105 | -35.0076 | -28.5346
18.4166 | -35.0081 | -28.5418
18.4227 | -35.0085 | -28.549
18.4288 | -35.009 | -28.5559
18.4349 | -35.0095 | -28.5629
18.441 | -35.0099 | -28.5702
18.4471 | -35.0104 | -28.5776
18.4532 | -35.0109 | -28.5846
18.4593 | -35.0113 | -28.5918
18.4654 | -35.0118 | -28.5986
18.4715 | -35.0123 | -28.606
18.4776 | -35.0128 | -28.6133
18.4837 | -35.0133 | -28.6203
18.4898 | -35.0137 | -28.6276
18.4959 | -35.0142 | -28.6349
18.502 | -35.0147 | -28.642
18.5081 | -35.0151 | -28.6492
18.5142 | -35.0156 | -28.6563
18.5203 | -35.0161 | -28.6637
18.5264 | -35.0166 | -28.6706
18.5325 | -35.0171 | -28.6778
18.5386 | -35.0176 | -28.6849
18.5447 | -35.0181 | -28.692
18.5508 | -35.0185 | -28.6989
18.557 | -35.019 | -28.7062
18.5631 | -35.0195 | -28.7135
18.5692 | -35.0199 | -28.7207
18.5753 | -35.0204 | -28.7276
18.5814 | -35.0209 | -28.7349
18.5875 | -35.0214 | -28.7421
18.5936 | -35.0218 | -28.7491
18.5997 | -35.0223 | -28.7562
18.6058 | -35.0227 | -28.7633
18.6119 | -35.0232 | -28.7703
18.618 | -35.0237 | -28.7778
18.6241 | -35.0242 | -28.7849
18.6302 | -35.0246 | -28.792
18.6363 | -35.0251 | -28.7992
18.6424 | -35.0256 | -28.8062
18.6485 | -35.0261 | -28.8133
18.6546 | -35.0266 | -28.8206
18.6607 | -35.0269 | -28.8276
18.6668 | -35.0274 | -28.8344
18.6729 | -35.0279 | -28.8417
18.679 | -35.0283 | -28.8485
18.6851 | -35.0288 | -28.8555
18.6912 | -35.0293 | -28.8624
18.6974 | -35.0298 | -28.8695
18.7035 | -35.0303 | -28.8765
18.7096 | -35.0308 | -28.8838
18.7157 | -35.0313 | -28.8908
18.7218 | -35.0317 | -28.8977
18.7279 | -35.0322 | -28.9049
18.734 | -35.0327 | -28.912
18.7401 | -35.0332 | -28.9192
18.7462 | -35.0337 | -28.9261
18.7523 | -35.0341 | -28.9331
18.7584 | -35.0346 | -28.9401
18.7645 | -35.0351 | -28.9471
18.7706 | -35.0356 | -28.954
18.7767 | -35.036 | -28.9612
18.7828 | -35.0365 | -28.9681
18.7889 | -35.037 | -28.9751
18.795 | -35.0375 | -28.9823
18.8011 | -35.0379 | -28.9893
18.8072 | -35.0384 | -28.9966
18.8133 | -35.0389 | -29.0038
18.8194 | -35.0394 | -29.0106
18.8255 | -35.0398 | -29.0177
18.8316 | -35.0403 | -29.0249
18.8377 | -35.0408 | -29.0319
18.8439 | -35.0413 | -29.0389
18.85 | -35.0418 | -29.0459
18.8561 | -35.0423 | -29.0527
18.8622 | -35.0427 | -29.0599
18.8683 | -35.0432 | -29.0668
18.8744 | -35.0437 | -29.0739
18.8805 | -35.0441 | -29.0813
18.8866 | -35.0446 | -29.0885
18.8927 | -35.045 | -29.0953
18.8988 | -35.0455 | -29.1025
18.9049 | -35.046 | -29.1096
18.911 | -35.0464 | -29.1167
18.9171 | -35.0469 | -29.1238
18.9232 | -35.0474 | -29.1307
18.9293 | -35.0479 | -29.1377
18.9354 | -35.0483 | -29.1446
18.9415 | -35.0488 | -29.1516
18.9476 | -35.0493 | -29.1586
18.9537 | -35.0498 | -29.1656
18.9598 | -35.0503 | -29.1724
18.9659 | -35.0507 | -29.1794
18.972 | -35.0512 | -29.1863
18.9781 | -35.0517 | -29.1933
18.9843 | -35.0521 | -29.2005
18.9904 | -35.0526 | -29.2076
18.9965 | -35.0531 | -29.2145
19.0026 | -35.0536 | -29.2214
19.0087 | -35.0541 | -29.2285
19.0148 | -35.0545 | -29.2356
19.0209 | -35.055 | -29.2427
19.027 | -35.0555 | -29.2497
19.0331 | -35.056 | -29.2564
19.0392 | -35.0565 | -29.2637
19.0453 | -35.057 | -29.2707
19.0514 | -35.0574 | -29.2777
19.0575 | -35.0579 | -29.2849
19.0636 | -35.0584 | -29.2916
19.0697 | -35.0589 | -29.2986
19.0758 | -35.0593 | -29.3056
19.0819 | -35.0598 | -29.3128
19.088 | -35.0603 | -29.3196
19.0941 | -35.0608 | -29.3267
19.1002 | -35.0612 | -29.3335
19.1063 | -35.0617 | -29.3407
19.1124 | -35.0622 | -29.3476
19.1185 | -35.0626 | -29.3546
19.1246 | -35.0631 | -29.3617
19.1308 | -35.0636 | -29.3689
19.1369 | -35.0641 | -29.3757
19.143 | -35.0646 | -29.3825
19.1491 | -35.0651 | -29.3894
19.1552 | -35.0655 | -29.3964
19.1613 | -35.066 | -29.4035
19.1674 | -35.0665 | -29.4107
19.1735 | -35.0669 | -29.4176
19.1796 | -35.0674 | -29.4246
19.1857 | -35.0679 | -29.4317
19.1918 | -35.0683 | -29.4389
19.1979 | -35.0688 | -29.4462
19.204 | -35.0693 | -29.4529
19.2101 | -35.0698 | -29.4597
19.2162 | -35.0703 | -29.4665
19.2223 | -35.0708 | -29.4735
19.2284 | -35.0712 | -29.4804
19.2345 | -35.0717 | -29.4874
19.2406 | -35.0722 | -29.4942
19.2467 | -35.0726 | -29.5012
19.2528 | -35.0731 | -29.5081
19.2589 | -35.0736 | -29.515
19.265 | -35.0741 | -29.522
19.2712 | -35.0745 | -29.5288
19.2773 | -35.075 | -29.5356
19.2834 | -35.0755 | -29.5427
19.2895 | -35.0759 | -29.55
19.2956 | -35.0764 | -29.5569
19.3017 | -35.0769 | -29.5636
19.3078 | -35.0773 | -29.5707
19.3139 | -35.0778 | -29.5773
19.32 | -35.0783 | -29.5841
19.3261 | -35.0787 | -29.5911
19.3322 | -35.0792 | -29.5982
19.3383 | -35.0797 | -29.6051
19.3444 | -35.0802 | -29.6118
19.3505 | -35.0807 | -29.6186
19.3566 | -35.0811 | -29.6255
19.3627 | -35.0816 | -29.6324
19.3688 | -35.0821 | -29.6396
19.3749 | -35.0826 | -29.6467
19.381 | -35.0831 | -29.6533
19.3871 | -35.0835 | -29.6604
19.3932 | -35.084 | -29.6672
19.3993 | -35.0845 | -29.6741
19.4054 | -35.0849 | -29.6811
19.4115 | -35.0854 | -29.6883
19.4177 | -35.0859 | -29.695
19.4238 | -35.0864 | -29.7017
19.4299 | -35.0868 | -29.7085
19.436 | -35.0873 | -29.7156
19.4421 | -35.0878 | -29.7224
19.4482 | -35.0883 | -29.7293
19.4543 | -35.0887 | -29.736
19.4604 | -35.0892 | -29.7429
19.4665 | -35.0896 | -29.7498
19.4726 | -35.0901 | -29.7565
19.4787 | -35.0905 | -29.7635
19.4848 | -35.091 | -29.7703
19.4909 | -35.0915 | -29.7771
19.497 | -35.092 | -29.7839
19.5031 | -35.0925 | -29.7909
19.5092 | -35.093 | -29.7977
19.5153 | -35.0934 | -29.8045
19.5214 | -35.0939 | -29.8111
19.5275 | -35.0944 | -29.8181
19.5336 | -35.0949 | -29.8251
19.5397 | -35.0953 | -29.832
19.5458 | -35.0958 | -29.8389
19.5519 | -35.0963 | -29.8456
19.5581 | -35.0967 | -29.8523
19.5642 | -35.0972 | -29.8592
19.5703 | -35.0976 | -29.866
19.5764 | -35.0981 | -29.8728
19.5825 | -35.0986 | -29.8795
19.5886 | -35.0991 | -29.8866
19.5947 | -35.0996 | -29.8934
19.6008 | -35.1001 | -29.9003
19.6069 | -35.1005 | -29.9074
19.613 | -35.101 | -29.9145
19.6191 | -35.1015 | -29.9211
19.6252 | -35.1019 | -29.9279
19.6313 | -35.1024 | -29.9343
19.6374 | -35.1029 | -29.9414
19.6435 | -35.1033 | -29.9483
19.6496 | -35.1038 | -29.9548
19.6557 | -35.1042 | -29.9616
19.6618 | -35.1047 | -29.9686
19.6679 | -35.1052 | -29.9755
19.674 | -35.1057 | -29.9823
19.6801 | -35.1061 | -29.9891
19.6862 | -35.1066 | -29.996
19.6923 | -35.1071 | -30.0029
19.6984 | -35.1076 | -30.0095
19.7046 | -35.1081 | -30.0165
19.7107 | -35.1086 | -30.0233
19.7168 | -35.109 | -30.0302
19.7229 | -35.1095 | -30.0372
19.729 | -35.11 | -30.044
19.7351 | -35.1105 | -30.0506
19.7412 | -35.111 | -30.0577
19.7473 | -35.1114 | -30.0643
19.7534 | -35.1119 | -30.0711
19.7595 | -35.1124 | -30.078
19.7656 | -35.1128 | -30.0848
19.7717 | -35.1133 | -30.0913
19.7778 | -35.1138 | -30.0984
19.7839 | -35.1142 | -30.1053
19.79 | -35.1147 | -30.1121
19.7961 | -35.1152 | -30.1191
19.8022 | -35.1156 | -30.1258
19.8083 | -35.1162 | -30.1326
19.8144 | -35.1167 | -30.1394
19.8205 | -35.1171 | -30.146
19.8266 | -35.1176 | -30.1524
19.8327 | -35.1181 | -30.1592
19.8388 | -35.1186 | -30.1655
19.845 | -35.119 | -30.1723
19.8511 | -35.1195 | -30.1791
19.8572 | -35.1199 | -30.186
19.8633 | -35.1204 | -30.1927
19.8694 | -35.1209 | -30.1995
19.8755 | -35.1214 | -30.2063
19.8816 | -35.1218 | -30.2133
19.8877 | -35.1223 | -30.2205
19.8938 | -35.1228 | -30.2269
19.8999 | -35.1233 | -30.2336
19.906 | -35.1238 | -30.2405
19.9121 | -35.1242 | -30.2471
19.9182 | -35.1247 | -30.2538
19.9243 | -35.1252 | -30.2603
19.9304 | -35.1256 | -30.2671
19.9365 | -35.1261 | -30.2738
19.9426 | -35.1266 | -30.2804
19.9487 | -35.1271 | -30.2872
19.9548 | -35.1276 | -30.2943
19.9609 | -35.1281 | -30.3013
19.967 | -35.1286 | -30.3079
19.9731 | -35.129 | -30.3149
19.9792 | -35.1295 | -30.3212
19.9853 | -35.1301 | -30.3277
19.9915 | -35.1306 | -30.3342
19.9976 | -35.1311 | -30.3411
20.0037 | -35.1316 | -30.3478
20.0098 | -35.1321 | -30.3549
20.0159 | -35.1326 | -30.3616
20.022 | -35.133 | -30.3686
20.0281 | -35.1335 | -30.3757
20.0342 | -35.134 | -30.3826
20.0403 | -35.1344 | -30.3891
20.0464 | -35.1349 | -30.3956
20.0525 | -35.1354 | -30.4023
20.0586 | -35.1359 | -30.4091
20.0647 | -35.1364 | -30.4159
20.0708 | -35.1369 | -30.4227
20.0769 | -35.1374 | -30.4297
20.083 | -35.1377 | -30.4358
20.0891 | -35.1382 | -30.4423
20.0952 | -35.1387 | -30.449
20.1013 | -35.1393 | -30.4557
20.1074 | -35.1397 | -30.4625
20.1135 | -35.1402 | -30.4692
20.1196 | -35.1407 | -30.4758
20.1257 | -35.1411 | -30.483
20.1319 | -35.1416 | -30.4899
20.138 | -35.1421 | -30.4969
20.1441 | -35.1427 | -30.5037
20.1502 | -35.1431 | -30.5105
20.1563 | -35.1436 | -30.5171
20.1624 | -35.1441 | -30.5237
20.1685 | -35.1446 | -30.5303
20.1746 | -35.1451 | -30.5373
20.1807 | -35.1455 | -30.5438
20.1868 | -35.146 | -30.5508
20.1929 | -35.1465 | -30.5571
20.199 | -35.147 | -30.5641
20.2051 | -35.1474 | -30.5704
20.2112 | -35.148 | -30.5769
20.2173 | -35.1484 | -30.5838
20.2234 | -35.1489 | -30.5904
20.2295 | -35.1494 | -30.5969
20.2356 | -35.1499 | -30.6039
20.2417 | -35.1504 | -30.61
20.2478 | -35.1509 | -30.6169
20.2539 | -35.1514 | -30.6236
20.26 | -35.1519 | -30.6305
20.2661 | -35.1523 | -30.6374
20.2723 | -35.1529 | -30.6438
20.2784 | -35.1533 | -30.6505
20.2845 | -35.1538 | -30.6571
20.2906 | -35.1543 | -30.6633
20.2967 | -35.1547 | -30.6703
20.3028 | -35.1552 | -30.6769
20.3089 | -35.1557 | -30.6839
20.315 | -35.1562 | -30.6907
20.3211 | -35.1567 | -30.6976
20.3272 | -35.1572 | -30.7043
20.3333 | -35.1577 | -30.7107
20.3394 | -35.1582 | -30.7172
20.3455 | -35.1586 | -30.7241
20.3516 | -35.159 | -30.7305
20.3577 | -35.1595 | -30.7373
20.3638 | -35.1599 | -30.7438
20.3699 | -35.1604 | -30.7503
20.376 | -35.1609 | -30.757
20.3821 | -35.1614 | -30.7637
20.3882 | -35.1619 | -30.7703
20.3943 | -35.1624 | -30.777
20.4004 | -35.1629 | -30.7834
20.4065 | -35.1634 | -30.7901
20.4126 | -35.1639 | -30.7964
20.4188 | -35.1644 | -30.803
20.4249 | -35.1649 | -30.8096
20.431 | -35.1654 | -30.8161
20.4371 | -35.1659 | -30.8225
20.4432 | -35.1664 | -30.8294
20.4493 | -35.1668 | -30.8361
20.4554 | -35.1673 | -30.8426
20.4615 | -35.1678 | -30.8494
20.4676 | -35.1683 | -30.8562
20.4737 | -35.1688 | -30.8627
20.4798 | -35.1693 | -30.8692
20.4859 | -35.1697 | -30.8759
20.492 | -35.1702 | -30.8826
20.4981 | -35.1707 | -30.8895
20.5042 | -35.1712 | -30.8959
20.5103 | -35.1716 | -30.9022
20.5164 | -35.1721 | -30.9087
20.5225 | -35.1726 | -30.9156
20.5286 | -35.173 | -30.9221
20.5347 | -35.1735 | -30.9289
20.5408 | -35.1739 | -30.9353
20.5469 | -35.1744 | -30.9419
20.553 | -35.1749 | -30.9488
20.5592 | -35.1754 | -30.9555
20.5653 | -35.1759 | -30.9621
20.5714 | -35.1764 | -30.9689
20.5775 | -35.1768 | -30.9748
20.5836 | -35.1773 | -30.9814
20.5897 | -35.1778 | -30.9881
20.5958 | -35.1783 | -30.9945
20.6019 | -35.1787 | -31.0011
20.608 | -35.1792 | -31.0077
20.6141 | -35.1796 | -31.0138
20.6202 | -35.1801 | -31.0202
20.6263 | -35.1806 | -31.0269
20.6324 | -35.1811 | -31.0334
20.6385 | -35.1816 | -31.0395
20.6446 | -35.1821 | -31.0461
20.6507 | -35.1826 | -31.0523
20.6568 | -35.183 | -31.0588
20.6629 | -35.1836 | -31.0656
20.669 | -35.1841 | -31.0721
20.6751 | -35.1846 | -31.0784
20.6812 | -35.1851 | -31.0846
20.6873 | -35.1856 | -31.091
20.6934 | -35.1861 | -31.0977
20.6995 | -35.1865 | -31.1041
20.7057 | -35.187 | -31.1107
20.7118 | -35.1876 | -31.1173
20.7179 | -35.1881 | -31.1234
20.724 | -35.1885 | -31.13
20.7301 | -35.189 | -31.1363
20.7362 | -35.1895 | -31.1429
20.7423 | -35.1899 | -31.1498
20.7484 | -35.1905 | -31.1567
20.7545 | -35.1909 | -31.1631
20.7606 | -35.1914 | -31.1698
20.7667 | -35.1919 | -31.1765
20.7728 | -35.1924 | -31.1831
20.7789 | -35.1929 | -31.19
20.785 | -35.1934 | -31.1965
20.7911 | -35.1939 | -31.2035
20.7972 | -35.1944 | -31.2101
20.8033 | -35.1949 | -31.2166
20.8094 | -35.1954 | -31.2229
20.8155 | -35.1959 | -31.2293
20.8216 | -35.1964 | -31.2354
20.8277 | -35.1968 | -31.242
20.8338 | -35.1972 | -31.2485
20.8399 | -35.1977 | -31.2549
20.8461 | -35.1982 | -31.2614
20.8522 | -35.1987 | -31.2683
20.8583 | -35.1993 | -31.2748
20.8644 | -35.1997 | -31.2812
20.8705 | -35.2001 | -31.2874
20.8766 | -35.2006 | -31.2939
20.8827 | -35.2011 | -31.3004
20.8888 | -35.2015 | -31.3072
20.8949 | -35.2021 | -31.3137
20.901 | -35.2025 | -31.3202
20.9071 | -35.203 | -31.3266
20.9132 | -35.2034 | -31.3329
20.9193 | -35.2038 | -31.3395
20.9254 | -35.2043 | -31.3463
20.9315 | -35.2049 | -31.3527
20.9376 | -35.2054 | -31.3594
20.9437 | -35.2058 | -31.3662
20.9498 | -35.2063 | -31.3728
20.9559 | -35.2068 | -31.3792
20.962 | -35.2073 | -31.3853
20.9681 | -35.2078 | -31.3918
20.9742 | -35.2083 | -31.3986
20.9803 | -35.2087 | -31.4051
20.9864 | -35.2092 | -31.4119
20.9926 | -35.2097 | -31.4184
20.9987 | -35.2101 | -31.4245
21.0048 | -35.2106 | -31.431
21.0109 | -35.211 | -31.4375
21.017 | -35.2116 | -31.4441
21.0231 | -35.212 | -31.4502
21.0292 | -35.2124 | -31.4567
21.0353 | -35.2129 | -31.4629
21.0414 | -35.2133 | -31.4693
21.0475 | -35.2139 | -31.4757
21.0536 | -35.2143 | -31.4823
21.0597 | -35.2148 | -31.4889
21.0658 | -35.2152 | -31.4955
21.0719 | -35.2157 | -31.5017
21.078 | -35.2161 | -31.5087
21.0841 | -35.2165 | -31.5148
21.0902 | -35.217 | -31.5212
21.0963 | -35.2175 | -31.5277
21.1024 | -35.218 | -31.5344
21.1085 | -35.2185 | -31.541
21.1146 | -35.219 | -31.5475
21.1207 | -35.2195 | -31.5545
21.1268 | -35.2199 | -31.561
21.133 | -35.2205 | -31.5678
21.1391 | -35.2209 | -31.5746
21.1452 | -35.2213 | -31.5806
21.1513 | -35.2218 | -31.5865
21.1574 | -35.2224 | -31.5931
21.1635 | -35.2228 | -31.5996
21.1696 | -35.2233 | -31.6061
21.1757 | -35.2238 | -31.6127
21.1818 | -35.2242 | -31.6196
21.1879 | -35.2246 | -31.6263
21.194 | -35.2252 | -31.6329
21.2001 | -35.2257 | -31.639
21.2062 | -35.2262 | -31.6458
21.2123 | -35.2266 | -31.6522
21.2184 | -35.2271 | -31.6579
21.2245 | -35.2275 | -31.6649
21.2306 | -35.228 | -31.6717
21.2367 | -35.2284 | -31.6781
21.2428 | -35.2289 | -31.6845
21.2489 | -35.2294 | -31.6912
21.255 | -35.2298 | -31.6982
21.2611 | -35.2303 | -31.7044
21.2672 | -35.2307 | -31.7112
21.2733 | -35.2312 | -31.7177
21.2795 | -35.2316 | -31.7241
21.2856 | -35.232 | -31.73
21.2917 | -35.2325 | -31.7366
21.2978 | -35.2329 | -31.7431
21.3039 | -35.2335 | -31.7494
21.31 | -35.234 | -31.7558
21.3161 | -35.2346 | -31.7625
21.3222 | -35.2351 | -31.7692
21.3283 | -35.2355 | -31.7753
21.3344 | -35.236 | -31.7814
21.3405 | -35.2364 | -31.7876
21.3466 | -35.2369 | -31.7943
21.3527 | -35.2374 | -31.8002
21.3588 | -35.2378 | -31.8066
21.3649 | -35.2382 | -31.8131
21.371 | -35.2387 | -31.8195
21.3771 | -35.2391 | -31.8256
21.3832 | -35.2396 | -31.8325
21.3893 | -35.2401 | -31.8386
21.3954 | -35.2406 | -31.8448
21.4015 | -35.2411 | -31.8514
21.4076 | -35.2416 | -31.8577
21.4137 | -35.242 | -31.8643
21.4199 | -35.2426 | -31.8707
21.426 | -35.243 | -31.877
21.4321 | -35.2435 | -31.8834
21.4382 | -35.2439 | -31.89
21.4443 | -35.2443 | -31.8962
21.4504 | -35.2448 | -31.9026
21.4565 | -35.2454 | -31.9091
21.4626 | -35.2459 | -31.9154
21.4687 | -35.2464 | -31.9216
21.4748 | -35.2468 | -31.9284
21.4809 | -35.2473 | -31.9344
21.487 | -35.2478 | -31.9402
21.4931 | -35.2482 | -31.9461
21.4992 | -35.2487 | -31.953
21.5053 | -35.2491 | -31.9596
21.5114 | -35.2495 | -31.9657
21.5175 | -35.25 | -31.9719
21.5236 | -35.2505 | -31.9782
21.5297 | -35.251 | -31.9849
21.5358 | -35.2515 | -31.991
21.5419 | -35.252 | -31.9973
21.548 | -35.2524 | -32.0037
21.5541 | -35.2527 | -32.0105
21.5602 | -35.2532 | -32.0169
21.5664 | -35.2538 | -32.023
21.5725 | -35.2543 | -32.0292
21.5786 | -35.2546 | -32.036
21.5847 | -35.2552 | -32.0424
21.5908 | -35.2557 | -32.0488
21.5969 | -35.2562 | -32.0549
21.603 | -35.2567 | -32.0613
21.6091 | -35.2572 | -32.0671
21.6152 | -35.2576 | -32.0735
21.6213 | -35.2581 | -32.08
21.6274 | -35.2586 | -32.0859
21.6335 | -35.259 | -32.092
21.6396 | -35.2595 | -32.0985
21.6457 | -35.2599 | -32.1046
21.6518 | -35.2604 | -32.1113
21.6579 | -35.261 | -32.1178
21.664 | -35.2614 | -32.1245
21.6701 | -35.2618 | -32.1305
21.6762 | -35.2622 | -32.1366
21.6823 | -35.2627 | -32.1431
21.6884 | -35.2632 | -32.1497
21.6945 | -35.2637 | -32.1564
21.7006 | -35.2641 | -32.163
21.7068 | -35.2646 | -32.169
21.7129 | -35.2651 | -32.1754
21.719 | -35.2656 | -32.1816
21.7251 | -35.266 | -32.1879
21.7312 | -35.2665 | -32.1945
21.7373 | -35.267 | -32.2005
21.7434 | -35.2674 | -32.2068
21.7495 | -35.2678 | -32.213
21.7556 | -35.2682 | -32.2197
21.7617 | -35.2687 | -32.2259
21.7678 | -35.2692 | -32.2326
21.7739 | -35.2697 | -32.239
21.78 | -35.2702 | -32.2455
21.7861 | -35.2707 | -32.2517
21.7922 | -35.2712 | -32.2582
21.7983 | -35.2716 | -32.2646
21.8044 | -35.2721 | -32.2712
21.8105 | -35.2726 | -32.2775
21.8166 | -35.273 | -32.2841
21.8227 | -35.2735 | -32.2906
21.8288 | -35.2739 | -32.2971
21.8349 | -35.2744 | -32.3036
21.841 | -35.2749 | -32.3099
21.8471 | -35.2754 | -32.3158
21.8533 | -35.2759 | -32.3218
21.8594 | -35.2763 | -32.3282
21.8655 | -35.2767 | -32.334
21.8716 | -35.2771 | -32.3403
21.8777 | -35.2775 | -32.346
21.8838 | -35.278 | -32.3528
21.8899 | -35.2784 | -32.3585
21.896 | -35.2788 | -32.3647
21.9021 | -35.2792 | -32.371
21.9082 | -35.2796 | -32.3777
21.9143 | -35.28 | -32.3839
21.9204 | -35.2804 | -32.3903
21.9265 | -35.2809 | -32.3966
21.9326 | -35.2814 | -32.4032
21.9387 | -35.2819 | -32.4092
21.9448 | -35.2824 | -32.4153
21.9509 | -35.2828 | -32.4217
21.957 | -35.2833 | -32.4278
21.9631 | -35.2837 | -32.4341
21.9692 | -35.2842 | -32.4404
21.9753 | -35.2846 | -32.4467
21.9814 | -35.2851 | -32.453
21.9875 | -35.2856 | -32.4593
21.9937 | -35.2861 | -32.4653
21.9998 | -35.2865 | -32.472
22.0059 | -35.2869 | -32.479
22.012 | -35.2874 | -32.4853
22.0181 | -35.2878 | -32.4911
22.0242 | -35.2883 | -32.4973
22.0303 | -35.2889 | -32.5038
22.0364 | -35.2893 | -32.5103
22.0425 | -35.2898 | -32.5165
22.0486 | -35.2903 | -32.5226
22.0547 | -35.2907 | -32.5285
22.0608 | -35.2912 | -32.5344
22.0669 | -35.2917 | -32.5403
22.073 | -35.2921 | -32.5464
22.0791 | -35.2926 | -32.5524
22.0852 | -35.293 | -32.5583
22.0913 | -35.2934 | -32.5643
22.0974 | -35.2939 | -32.5709
22.1035 | -35.2944 | -32.5773
22.1096 | -35.2948 | -32.5831
22.1157 | -35.2953 | -32.5897
22.1218 | -35.2957 | -32.5957
22.1279 | -35.2962 | -32.6013
22.134 | -35.2967 | -32.6081
22.1402 | -35.2972 | -32.6141
22.1463 | -35.2977 | -32.6207
22.1524 | -35.2981 | -32.6269
22.1585 | -35.2986 | -32.6333
22.1646 | -35.299 | -32.641
22.1707 | -35.2994 | -32.6472
22.1768 | -35.2999 | -32.6542
22.1829 | -35.3003 | -32.6605
22.189 | -35.3007 | -32.667
22.1951 | -35.3011 | -32.6727
22.2012 | -35.3015 | -32.6797
22.2073 | -35.302 | -32.6858
22.2134 | -35.3024 | -32.6919
22.2195 | -35.3029 | -32.6985
22.2256 | -35.3034 | -32.7048
22.2317 | -35.3037 | -32.7108
22.2378 | -35.3041 | -32.7166
22.2439 | -35.3045 | -32.7229
22.25 | -35.3049 | -32.7296
22.2561 | -35.3054 | -32.7359
22.2622 | -35.3058 | -32.7421
22.2683 | -35.3061 | -32.7486
22.2744 | -35.3066 | -32.7547
22.2806 | -35.307 | -32.7614
22.2867 | -35.3075 | -32.7678
22.2928 | -35.308 | -32.7739
22.2989 | -35.3085 | -32.7801
22.305 | -35.309 | -32.7861
22.3111 | -35.3094 | -32.7924
22.3172 | -35.3099 | -32.7991
22.3233 | -35.3104 | -32.8051
22.3294 | -35.3108 | -32.8114
22.3355 | -35.3112 | -32.8176
22.3416 | -35.3116 | -32.8239
22.3477 | -35.3121 | -32.8302
22.3538 | -35.3125 | -32.8363
22.3599 | -35.3129 | -32.8427
22.366 | -35.3134 | -32.8487
22.3721 | -35.314 | -32.854
22.3782 | -35.3143 | -32.8608
22.3843 | -35.3147 | -32.8673
22.3904 | -35.3152 | -32.8737
22.3965 | -35.3155 | -32.8802
22.4026 | -35.3161 | -32.8866
22.4087 | -35.3165 | -32.8925
22.4148 | -35.3169 | -32.8984
22.4209 | -35.3173 | -32.9049
22.4271 | -35.3177 | -32.9109
22.4332 | -35.3181 | -32.917
22.4393 | -35.3185 | -32.9234
22.4454 | -35.319 | -32.9296
22.4515 | -35.3194 | -32.9356
22.4576 | -35.3198 | -32.9419
22.4637 | -35.3203 | -32.9475
22.4698 | -35.3207 | -32.9538
22.4759 | -35.3212 | -32.9598
22.482 | -35.3216 | -32.9664
22.4881 | -35.322 | -32.9727
22.4942 | -35.3224 | -32.9792
22.5003 | -35.323 | -32.9854
22.5064 | -35.3234 | -32.9921
22.5125 | -35.3238 | -32.9978
22.5186 | -35.3242 | -33.0043
22.5247 | -35.3247 | -33.0103
22.5308 | -35.3251 | -33.0165
22.5369 | -35.3255 | -33.023
22.543 | -35.3259 | -33.0294
22.5491 | -35.3263 | -33.0352
22.5552 | -35.3268 | -33.042
22.5613 | -35.3273 | -33.0485
22.5675 | -35.3277 | -33.0543
22.5736 | -35.3282 | -33.0606
22.5797 | -35.3286 | -33.0666
22.5858 | -35.329 | -33.0727
22.5919 | -35.3294 | -33.0788
22.598 | -35.3297 | -33.0852
22.6041 | -35.3301 | -33.0912
22.6102 | -35.3306 | -33.0976
22.6163 | -35.331 | -33.1034
22.6224 | -35.3315 | -33.1099
22.6285 | -35.3319 | -33.1153
22.6346 | -35.3324 | -33.1219
22.6407 | -35.3328 | -33.1282
22.6468 | -35.3333 | -33.135
22.6529 | -35.3336 | -33.1406
22.659 | -35.3339 | -33.1468
22.6651 | -35.3343 | -33.1526
22.6712 | -35.3347 | -33.1585
22.6773 | -35.3352 | -33.165
22.6834 | -35.3356 | -33.1706
22.6895 | -35.3359 | -33.1769
22.6956 | -35.3363 | -33.1826
22.7017 | -35.3367 | -33.1882
22.7079 | -35.3371 | -33.1941
22.714 | -35.3376 | -33.2003
22.7201 | -35.338 | -33.2062
22.7262 | -35.3385 | -33.2135
22.7323 | -35.3389 | -33.2199
22.7384 | -35.3393 | -33.2261
22.7445 | -35.3397 | -33.2324
22.7506 | -35.3402 | -33.2386
22.7567 | -35.3406 | -33.2451
22.7628 | -35.3412 | -33.2508
22.7689 | -35.3415 | -33.2572
22.775 | -35.342 | -33.2632
22.7811 | -35.3424 | -33.2696
22.7872 | -35.3428 | -33.2755
22.7933 | -35.3432 | -33.2818
22.7994 | -35.3435 | -33.2878
22.8055 | -35.344 | -33.2944
22.8116 | -35.3444 | -33.3006
22.8177 | -35.3449 | -33.3069
22.8238 | -35.3453 | -33.3129
22.8299 | -35.3458 | -33.3193
22.836 | -35.3461 | -33.3255
22.8421 | -35.3466 | -33.3318
22.8482 | -35.3469 | -33.3382
22.8544 | -35.3474 | -33.3444
22.8605 | -35.3478 | -33.3501
22.8666 | -35.3482 | -33.3561
22.8727 | -35.3486 | -33.3623
22.8788 | -35.349 | -33.3686
22.8849 | -35.3495 | -33.3748
22.891 | -35.3499 | -33.3811
22.8971 | -35.3503 | -33.3876
22.9032 | -35.3507 | -33.3933
22.9093 | -35.3512 | -33.3994
22.9154 | -35.3516 | -33.4052
22.9215 | -35.352 | -33.4113
22.9276 | -35.3525 | -33.418
22.9337 | -35.3529 | -33.424
22.9398 | -35.3534 | -33.4307
22.9459 | -35.3538 | -33.4371
22.952 | -35.3541 | -33.4437
22.9581 | -35.3545 | -33.4496
22.9642 | -35.3549 | -33.4563
22.9703 | -35.3553 | -33.4623
22.9764 | -35.3557 | -33.4683
22.9825 | -35.3561 | -33.4742
22.9886 | -35.3565 | -33.4807
22.9948 | -35.3569 | -33.4866
23.0009 | -35.3573 | -33.4928
23.007 | -35.3577 | -33.499
23.0131 | -35.358 | -33.505
23.0192 | -35.3585 | -33.5119
23.0253 | -35.3588 | -33.5177
23.0314 | -35.3593 | -33.5236
23.0375 | -35.3597 | -33.5297
23.0436 | -35.3602 | -33.5353
23.0497 | -35.3606 | -33.5422
23.0558 | -35.361 | -33.5477
23.0619 | -35.3613 | -33.5534
23.068 | -35.3618 | -33.5599
23.0741 | -35.3621 | -33.5658
23.0802 | -35.3625 | -33.5724
23.0863 | -35.363 | -33.5781
23.0924 | -35.3635 | -33.5842
23.0985 | -35.3639 | -33.5906
23.1046 | -35.3644 | -33.5964
23.1107 | -35.3647 | -33.6022
23.1168 | -35.3652 | -33.6085
23.1229 | -35.3655 | -33.6144
23.129 | -35.366 | -33.6204
23.1351 | -35.3664 | -33.6266
23.1413 | -35.3667 | -33.6327
23.1474 | -35.367 | -33.6395
23.1535 | -35.3674 | -33.6457
23.1596 | -35.3678 | -33.6521
23.1657 | -35.3683 | -33.6582
23.1718 | -35.3686 | -33.6641
23.1779 | -35.369 | -33.6699
23.184 | -35.3695 | -33.6754
23.1901 | -35.37 | -33.6819
23.1962 | -35.3705 | -33.6875
23.2023 | -35.371 | -33.6939
23.2084 | -35.3713 | -33.7001
23.2145 | -35.3716 | -33.7053
23.2206 | -35.372 | -33.7117
23.2267 | -35.3724 | -33.7173
23.2328 | -35.3728 | -33.7231
23.2389 | -35.3732 | -33.7294
23.245 | -35.3735 | -33.7357
23.2511 | -35.374 | -33.7409
23.2572 | -35.3745 | -33.7471
23.2633 | -35.3749 | -33.753
23.2694 | -35.3753 | -33.759
23.2755 | -35.3758 | -33.765
23.2817 | -35.3762 | -33.7712
23.2878 | -35.3766 | -33.7768
23.2939 | -35.3771 | -33.7835
23.3 | -35.3775 | -33.7907
23.3061 | -35.3778 | -33.796
23.3122 | -35.3782 | -33.8017
23.3183 | -35.3785 | -33.8081
23.3244 | -35.3789 | -33.8139
23.3305 | -35.3793 | -33.8199
23.3366 | -35.3798 | -33.8256
23.3427 | -35.3803 | -33.8313
23.3488 | -35.3808 | -33.8377
23.3549 | -35.3812 | -33.8442
23.361 | -35.3815 | -33.8501
23.3671 | -35.3819 | -33.8561
23.3732 | -35.3823 | -33.8619
23.3793 | -35.3827 | -33.8684
23.3854 | -35.3831 | -33.8744
23.3915 | -35.3836 | -33.8804
23.3976 | -35.384 | -33.886
23.4037 | -35.3844 | -33.8921
23.4098 | -35.3849 | -33.8985
23.4159 | -35.3853 | -33.9047
23.422 | -35.3857 | -33.9108
23.4282 | -35.3861 | -33.9168
23.4343 | -35.3866 | -33.9225
23.4404 | -35.387 | -33.9289
23.4465 | -35.3875 | -33.9353
23.4526 | -35.3879 | -33.9414
23.4587 | -35.3883 | -33.947
23.4648 | -35.3886 | -33.9525
23.4709 | -35.389 | -33.9582
23.477 | -35.3893 | -33.9641
23.4831 | -35.3896 | -33.9698
23.4892 | -35.39 | -33.9755
23.4953 | -35.3904 | -33.9822
23.5014 | -35.3908 | -33.988
23.5075 | -35.3912 | -33.9946
23.5136 | -35.3917 | -34.0009
23.5197 | -35.3921 | -34.0076
23.5258 | -35.3925 | -34.0132
23.5319 | -35.3929 | -34.0187
23.538 | -35.3933 | -34.0247
23.5441 | -35.3937 | -34.0306
23.5502 | -35.3941 | -34.0365
23.5563 | -35.3946 | -34.0426
23.5624 | -35.395 | -34.0486
23.5686 | -35.3955 | -34.0538
23.5747 | -35.3959 | -34.0602
23.5808 | -35.3962 | -34.0664
23.5869 | -35.3966 | -34.0722
23.593 | -35.397 | -34.0785
23.5991 | -35.3975 | -34.0838
23.6052 | -35.3979 | -34.0889
23.6113 | -35.3983 | -34.0952
23.6174 | -35.3987 | -34.1017
23.6235 | -35.3992 | -34.1076
23.6296 | -35.3996 | -34.1133
23.6357 | -35.3999 | -34.1186
23.6418 | -35.4005 | -34.1239
23.6479 | -35.4009 | -34.1301
23.654 | -35.4013 | -34.1361
23.6601 | -35.4017 | -34.1421
23.6662 | -35.4021 | -34.148
23.6723 | -35.4025 | -34.1539
23.6784 | -35.4029 | -34.16
23.6845 | -35.4033 | -34.1656
23.6906 | -35.4037 | -34.1718
23.6967 | -35.404 | -34.1779
23.7028 | -35.4045 | -34.1839
23.7089 | -35.4048 | -34.1897
23.7151 | -35.4053 | -34.1958
23.7212 | -35.4056 | -34.2016
23.7273 | -35.406 | -34.2078
23.7334 | -35.4064 | -34.2135
23.7395 | -35.4068 | -34.2197
23.7456 | -35.4073 | -34.2251
23.7517 | -35.4077 | -34.2307
23.7578 | -35.4081 | -34.2367
23.7639 | -35.4085 | -34.2427
23.77 | -35.409 | -34.2487
23.7761 | -35.4094 | -34.2547
23.7822 | -35.4098 | -34.2611
23.7883 | -35.4103 | -34.267
23.7944 | -35.4107 | -34.2731
23.8005 | -35.4112 | -34.2791
23.8066 | -35.4116 | -34.2846
23.8127 | -35.4118 | -34.2899
23.8188 | -35.4123 | -34.2964
23.8249 | -35.4126 | -34.3017
23.831 | -35.4129 | -34.307
23.8371 | -35.4132 | -34.3126
23.8432 | -35.4137 | -34.3181
23.8493 | -35.4141 | -34.3227
23.8555 | -35.4146 | -34.329
23.8616 | -35.4151 | -34.3352
23.8677 | -35.4154 | -34.3408
23.8738 | -35.4158 | -34.3473
23.8799 | -35.4162 | -34.3533
23.886 | -35.4165 | -34.3594
23.8921 | -35.4169 | -34.3649
23.8982 | -35.4173 | -34.3714
23.9043 | -35.4176 | -34.3776
23.9104 | -35.4179 | -34.3836
23.9165 | -35.4182 | -34.3898
23.9226 | -35.4186 | -34.396
23.9287 | -35.419 | -34.4019
23.9348 | -35.4194 | -34.4081
23.9409 | -35.4199 | -34.4139
23.947 | -35.4203 | -34.4204
23.9531 | -35.4208 | -34.4258
23.9592 | -35.4212 | -34.4313
23.9653 | -35.4216 | -34.4373
23.9714 | -35.4219 | -34.4433
23.9775 | -35.4223 | -34.4493
23.9836 | -35.4227 | -34.4554
23.9897 | -35.4231 | -34.4617
23.9958 | -35.4235 | -34.4677
24.002 | -35.4239 | -34.4739
24.0081 | -35.4243 | -34.4797
24.0142 | -35.4247 | -34.4855
24.0203 | -35.4252 | -34.4912
24.0264 | -35.4255 | -34.4968
24.0325 | -35.4259 | -34.5025
24.0386 | -35.4262 | -34.5089
24.0447 | -35.4265 | -34.5147
24.0508 | -35.4269 | -34.5207
24.0569 | -35.4273 | -34.5263
24.063 | -35.4277 | -34.5316
24.0691 | -35.4281 | -34.5373
24.0752 | -35.4285 | -34.543
24.0813 | -35.429 | -34.5493
24.0874 | -35.4294 | -34.5555
24.0935 | -35.4299 | -34.5615
24.0996 | -35.4302 | -34.5672
24.1057 | -35.4306 | -34.5733
24.1118 | -35.4311 | -34.579
24.1179 | -35.4315 | -34.5845
24.124 | -35.4319 | -34.5904
24.1301 | -35.4323 | -34.5955
24.1362 | -35.4327 | -34.601
24.1424 | -35.433 | -34.607
24.1485 | -35.4333 | -34.6131
24.1546 | -35.4338 | -34.6182
24.1607 | -35.4342 | -34.6238
24.1668 | -35.4345 | -34.6291
24.1729 | -35.4349 | -34.6346
24.179 | -35.4353 | -34.6408
24.1851 | -35.4358 | -34.6468
24.1912 | -35.4364 | -34.6524
24.1973 | -35.4365 | -34.6578
24.2034 | -35.437 | -34.664
24.2095 | -35.4374 | -34.6699
24.2156 | -35.4377 | -34.6761
24.2217 | -35.4381 | -34.6811
24.2278 | -35.4384 | -34.6874
24.2339 | -35.4387 | -34.6933
24.24 | -35.4392 | -34.6986
24.2461 | -35.4395 | -34.7045
24.2522 | -35.4399 | -34.7104
24.2583 | -35.4404 | -34.7163
24.2644 | -35.4407 | -34.7221
24.2705 | -35.4411 | -34.7278
24.2766 | -35.4415 | -34.7335
24.2827 | -35.4419 | -34.7388
24.2889 | -35.4423 | -34.7454
24.295 | -35.4425 | -34.7519
24.3011 | -35.4429 | -34.7579
24.3072 | -35.4433 | -34.7629
24.3133 | -35.4436 | -34.7692
24.3194 | -35.4441 | -34.7743
24.3255 | -35.4444 | -34.78
24.3316 | -35.4448 | -34.7857
24.3377 | -35.4452 | -34.791
24.3438 | -35.4455 | -34.7963
24.3499 | -35.4459 | -34.8023
24.356 | -35.4463 | -34.8082
24.3621 | -35.4467 | -34.8142
24.3682 | -35.4471 | -34.8204
24.3743 | -35.4473 | -34.8261
24.3804 | -35.4477 | -34.8314
24.3865 | -35.4482 | -34.8375
24.3926 | -35.4486 | -34.8431
24.3987 | -35.4489 | -34.8493
24.4048 | -35.4493 | -34.8547
24.4109 | -35.4497 | -34.8604
24.417 | -35.4501 | -34.8663
24.4231 | -35.4505 | -34.8723
24.4293 | -35.4508 | -34.8777
24.4354 | -35.4512 | -34.8834
24.4415 | -35.4516 | -34.8883
24.4476 | -35.452 | -34.8932
24.4537 | -35.4525 | -34.8993
24.4598 | -35.4528 | -34.9058
24.4659 | -35.4533 | -34.9121
24.472 | -35.4537 | -34.919
24.4781 | -35.454 | -34.9242
24.4842 | -35.4544 | -34.9293
24.4903 | -35.4548 | -34.9353
24.4964 | -35.4552 | -34.9409
24.5025 | -35.4556 | -34.9464
24.5086 | -35.456 | -34.9517
24.5147 | -35.4564 | -34.9578
24.5208 | -35.4567 | -34.9634
24.5269 | -35.4571 | -34.969
24.533 | -35.4576 | -34.9752
24.5391 | -35.458 | -34.9811
24.5452 | -35.4583 | -34.987
24.5513 | -35.4589 | -34.9925
24.5574 | -35.4593 | -34.9977
24.5635 | -35.4597 | -35.0038
24.5696 | -35.46 | -35.0096
24.5758 | -35.4604 | -35.0153
24.5819 | -35.4607 | -35.0213
24.588 | -35.4611 | -35.0267
24.5941 | -35.4615 | -35.0327
24.6002 | -35.4619 | -35.0383
24.6063 | -35.4624 | -35.0438
24.6124 | -35.4629 | -35.0498
24.6185 | -35.4633 | -35.0556
24.6246 | -35.4637 | -35.061
24.6307 | -35.4641 | -35.0674
24.6368 | -35.4645 | -35.0731
24.6429 | -35.4649 | -35.0789
24.649 | -35.4653 | -35.0849
24.6551 | -35.4656 | -35.0906
24.6612 | -35.4661 | -35.0965
24.6673 | -35.4664 | -35.1022
24.6734 | -35.4667 | -35.108
24.6795 | -35.4671 | -35.114
24.6856 | -35.4675 | -35.1197
24.6917 | -35.468 | -35.1256
24.6978 | -35.4684 | -35.1316
24.7039 | -35.4688 | -35.1378
24.71 | -35.4691 | -35.1439
24.7162 | -35.4695 | -35.1497
24.7223 | -35.4698 | -35.1551
24.7284 | -35.4701 | -35.1597
24.7345 | -35.4705 | -35.165
24.7406 | -35.471 | -35.1705
24.7467 | -35.4713 | -35.1772
24.7528 | -35.4716 | -35.183
24.7589 | -35.4719 | -35.1893
24.765 | -35.4724 | -35.1941
24.7711 | -35.4728 | -35.1994
24.7772 | -35.4732 | -35.205
24.7833 | -35.4734 | -35.2106
24.7894 | -35.4739 | -35.2169
24.7955 | -35.4743 | -35.2232
24.8016 | -35.4747 | -35.2296
24.8077 | -35.475 | -35.2351
24.8138 | -35.4755 | -35.2412
24.8199 | -35.4759 | -35.2473
24.826 | -35.4763 | -35.2536
24.8321 | -35.4767 | -35.2589
24.8382 | -35.4772 | -35.2643
24.8443 | -35.4776 | -35.2702
24.8504 | -35.478 | -35.2762
24.8565 | -35.4784 | -35.2813
24.8627 | -35.4787 | -35.2868
24.8688 | -35.4792 | -35.2931
24.8749 | -35.4796 | -35.2987
24.881 | -35.48 | -35.3042
24.8871 | -35.4805 | -35.3101
24.8932 | -35.4809 | -35.3157
24.8993 | -35.4813 | -35.3211
24.9054 | -35.4817 | -35.3277
24.9115 | -35.4821 | -35.333
24.9176 | -35.4824 | -35.3391
24.9237 | -35.4829 | -35.3449
24.9298 | -35.4833 | -35.35
24.9359 | -35.4836 | -35.356
24.942 | -35.4841 | -35.3613
24.9481 | -35.4846 | -35.3673
24.9542 | -35.485 | -35.3731
24.9603 | -35.4854 | -35.3789
24.9664 | -35.4857 | -35.3854
24.9725 | -35.4861 | -35.3907
24.9786 | -35.4867 | -35.3952
24.9847 | -35.4871 | -35.4005
24.9908 | -35.4875 | -35.4057
24.9969 | -35.4879 | -35.412
25.0031 | -35.4883 | -35.4177
25.0092 | -35.4887 | -35.4233
25.0153 | -35.4891 | -35.4291
25.0214 | -35.4894 | -35.4347
25.0275 | -35.49 | -35.4412
25.0336 | -35.4904 | -35.4474
25.0397 | -35.4908 | -35.4535
25.0458 | -35.4913 | -35.4578
25.0519 | -35.4917 | -35.4631
25.058 | -35.4921 | -35.469
25.0641 | -35.4925 | -35.474
25.0702 | -35.4927 | -35.4796
25.0763 | -35.4931 | -35.4852
25.0824 | -35.4936 | -35.4915
25.0885 | -35.4939 | -35.4981
25.0946 | -35.4942 | -35.5027
25.1007 | -35.4947 | -35.5081
25.1068 | -35.4951 | -35.5141
25.1129 | -35.4955 | -35.5195
25.119 | -35.496 | -35.5248
25.1251 | -35.4964 | -35.5304
25.1312 | -35.4968 | -35.5366
25.1373 | -35.497 | -35.5428
25.1435 | -35.4974 | -35.5485
25.1496 | -35.4977 | -35.5547
25.1557 | -35.498 | -35.5605
25.1618 | -35.4984 | -35.5661
25.1679 | -35.4988 | -35.5712
25.174 | -35.4992 | -35.5768
25.1801 | -35.4996 | -35.5824
25.1862 | -35.5 | -35.5874
25.1923 | -35.5003 | -35.593
25.1984 | -35.5007 | -35.599
25.2045 | -35.501 | -35.605
25.2106 | -35.5015 | -35.6118
25.2167 | -35.5019 | -35.6179
25.2228 | -35.5022 | -35.6233
25.2289 | -35.5026 | -35.6286
25.235 | -35.503 | -35.6341
25.2411 | -35.5033 | -35.6392
25.2472 | -35.5037 | -35.6447
25.2533 | -35.5041 | -35.6501
25.2594 | -35.5046 | -35.6556
25.2655 | -35.5049 | -35.6608
25.2716 | -35.5054 | -35.6661
25.2777 | -35.5058 | -35.6722
25.2838 | -35.5063 | -35.6776
25.29 | -35.5066 | -35.6837
25.2961 | -35.507 | -35.6888
25.3022 | -35.5074 | -35.6937
25.3083 | -35.5078 | -35.6994
25.3144 | -35.5082 | -35.7055
25.3205 | -35.5086 | -35.7101
25.3266 | -35.5091 | -35.7155
25.3327 | -35.5093 | -35.7212
25.3388 | -35.5097 | -35.7283
25.3449 | -35.5102 | -35.7346
25.351 | -35.5107 | -35.7406
25.3571 | -35.511 | -35.746
25.3632 | -35.5116 | -35.7522
25.3693 | -35.5119 | -35.7567
25.3754 | -35.5124 | -35.763
25.3815 | -35.5127 | -35.7685
25.3876 | -35.5132 | -35.7741
25.3937 | -35.5135 | -35.7798
25.3998 | -35.514 | -35.7851
25.4059 | -35.5145 | -35.7909
25.412 | -35.5149 | -35.7968
25.4181 | -35.5152 | -35.8025
25.4242 | -35.5156 | -35.808
25.4304 | -35.516 | -35.8133
25.4365 | -35.5165 | -35.8186
25.4426 | -35.5169 | -35.8249
25.4487 | -35.5174 | -35.8304
25.4548 | -35.5177 | -35.8357
25.4609 | -35.5181 | -35.8411
25.467 | -35.5184 | -35.8466
25.4731 | -35.5189 | -35.8527
25.4792 | -35.5193 | -35.8585
25.4853 | -35.5197 | -35.864
25.4914 | -35.5201 | -35.8696
25.4975 | -35.5205 | -35.8748
25.5036 | -35.521 | -35.8798
25.5097 | -35.5213 | -35.8855
25.5158 | -35.5217 | -35.8912
25.5219 | -35.522 | -35.8966
25.528 | -35.5225 | -35.9024
25.5341 | -35.5229 | -35.9079
25.5402 | -35.5234 | -35.9134
25.5463 | -35.5238 | -35.9193
25.5524 | -35.5241 | -35.9248
25.5585 | -35.5245 | -35.9302
25.5646 | -35.5248 | -35.936
25.5707 | -35.5253 | -35.942
25.5769 | -35.5257 | -35.9474
25.583 | -35.5261 | -35.953
25.5891 | -35.5265 | -35.9587
25.5952 | -35.5269 | -35.9645
25.6013 | -35.5273 | -35.9698
25.6074 | -35.5277 | -35.9746
25.6135 | -35.528 | -35.98
25.6196 | -35.5284 | -35.9855
25.6257 | -35.5287 | -35.9918
25.6318 | -35.5291 | -35.9971
25.6379 | -35.5295 | -36.0032
25.644 | -35.5299 | -36.0094
25.6501 | -35.5302 | -36.0159
25.6562 | -35.5307 | -36.0217
25.6623 | -35.531 | -36.0269
25.6684 | -35.5314 | -36.0326
25.6745 | -35.5317 | -36.0384
25.6806 | -35.5321 | -36.0435
25.6867 | -35.5324 | -36.0489
25.6928 | -35.5328 | -36.0543
25.6989 | -35.5333 | -36.06
25.705 | -35.5337 | -36.0658
25.7111 | -35.5342 | -36.0718
25.7173 | -35.5346 | -36.077
25.7234 | -35.535 | -36.083
25.7295 | -35.5354 | -36.0889
25.7356 | -35.5358 | -36.0942
25.7417 | -35.5362 | -36.0994
25.7478 | -35.5367 | -36.1052
25.7539 | -35.5371 | -36.1109
25.76 | -35.5374 | -36.1161
25.7661 | -35.5379 | -36.1212
25.7722 | -35.5384 | -36.1266
25.7783 | -35.5388 | -36.1323
25.7844 | -35.5392 | -36.1375
25.7905 | -35.5396 | -36.1435
25.7966 | -35.54 | -36.1491
25.8027 | -35.5404 | -36.155
25.8088 | -35.5408 | -36.1607
25.8149 | -35.5411 | -36.1663
25.821 | -35.5416 | -36.1717
25.8271 | -35.542 | -36.177
25.8332 | -35.5424 | -36.1831
25.8393 | -35.5427 | -36.1893
25.8454 | -35.5431 | -36.1952
25.8515 | -35.5435 | -36.2008
25.8576 | -35.5439 | -36.2064
25.8638 | -35.5445 | -36.2122
25.8699 | -35.5449 | -36.2174
25.876 | -35.5453 | -36.2223
25.8821 | -35.5457 | -36.2278
25.8882 | -35.5461 | -36.2336
25.8943 | -35.5465 | -36.2394
25.9004 | -35.5468 | -36.2441
25.9065 | -35.5473 | -36.2496
25.9126 | -35.5476 | -36.2551
25.9187 | -35.548 | -36.2602
25.9248 | -35.5484 | -36.2661
25.9309 | -35.549 | -36.2717
25.937 | -35.5493 | -36.2771
25.9431 | -35.5497 | -36.2822
25.9492 | -35.55 | -36.2878
25.9553 | -35.5504 | -36.293
25.9614 | -35.5507 | -36.2982
25.9675 | -35.5511 | -36.3036
25.9736 | -35.5515 | -36.3099
25.9797 | -35.5519 | -36.3155
25.9858 | -35.5522 | -36.3215
25.9919 | -35.5526 | -36.3267
25.998 | -35.5532 | -36.3323
26.0042 | -35.5535 | -36.3381
26.0103 | -35.5539 | -36.3438
26.0164 | -35.5544 | -36.3496
26.0225 | -35.5549 | -36.355
26.0286 | -35.5552 | -36.3607
26.0347 | -35.5556 | -36.3661
26.0408 | -35.5561 | -36.3718
26.0469 | -35.5565 | -36.3772
26.053 | -35.5568 | -36.3832
26.0591 | -35.5572 | -36.3886
26.0652 | -35.5576 | -36.3942
26.0713 | -35.558 | -36.3994
26.0774 | -35.5584 | -36.4048
26.0835 | -35.5588 | -36.4105
26.0896 | -35.5592 | -36.4159
26.0957 | -35.5596 | -36.4215
26.1018 | -35.56 | -36.427
26.1079 | -35.5603 | -36.4323
26.114 | -35.5608 | -36.4377
26.1201 | -35.5611 | -36.4431
26.1262 | -35.5615 | -36.4487
26.1323 | -35.5619 | -36.4538
26.1384 | -35.5622 | -36.4592
26.1445 | -35.5626 | -36.4646
26.1507 | -35.5631 | -36.4705
26.1568 | -35.5635 | -36.4759
26.1629 | -35.5639 | -36.4812
26.169 | -35.5643 | -36.4869
26.1751 | -35.5648 | -36.4919
26.1812 | -35.5652 | -36.4977
26.1873 | -35.5656 | -36.503
26.1934 | -35.5661 | -36.5084
26.1995 | -35.5666 | -36.5136
26.2056 | -35.5671 | -36.5191
26.2117 | -35.5675 | -36.5243
26.2178 | -35.5679 | -36.5297
26.2239 | -35.5684 | -36.5359
26.23 | -35.5687 | -36.5413
26.2361 | -35.569 | -36.547
26.2422 | -35.5692 | -36.5529
26.2483 | -35.5697 | -36.5577
26.2544 | -35.5701 | -36.5631
26.2605 | -35.5706 | -36.5693
26.2666 | -35.571 | -36.5755
26.2727 | -35.5714 | -36.5807
26.2788 | -35.5718 | -36.586
26.2849 | -35.5723 | -36.5916
26.2911 | -35.5727 | -36.5973
26.2972 | -35.573 | -36.6028
26.3033 | -35.5734 | -36.6083
26.3094 | -35.5738 | -36.6133
26.3155 | -35.5742 | -36.6186
26.3216 | -35.5745 | -36.6244
26.3277 | -35.5748 | -36.6299
26.3338 | -35.5752 | -36.6352
26.3399 | -35.5756 | -36.6406
26.346 | -35.5762 | -36.6456
26.3521 | -35.5766 | -36.6518
26.3582 | -35.5769 | -36.6576
26.3643 | -35.5773 | -36.6629
26.3704 | -35.5778 | -36.668
26.3765 | -35.5783 | -36.673
26.3826 | -35.5787 | -36.6785
26.3887 | -35.5791 | -36.6843
26.3948 | -35.5794 | -36.6902
26.4009 | -35.5798 | -36.6956
26.407 | -35.5802 | -36.7011
26.4131 | -35.5806 | -36.7065
26.4192 | -35.581 | -36.7119
26.4253 | -35.5814 | -36.7176
26.4314 | -35.5818 | -36.7235
26.4376 | -35.5823 | -36.7284
26.4437 | -35.5826 | -36.7338
26.4498 | -35.5831 | -36.7394
26.4559 | -35.5835 | -36.7451
26.462 | -35.5839 | -36.7508
26.4681 | -35.5843 | -36.756
26.4742 | -35.5847 | -36.7613
26.4803 | -35.5852 | -36.7666
26.4864 | -35.5856 | -36.7715
26.4925 | -35.586 | -36.7767
26.4986 | -35.5863 | -36.7818
26.5047 | -35.5868 | -36.7874
26.5108 | -35.5872 | -36.7927
26.5169 | -35.5876 | -36.7981
26.523 | -35.588 | -36.8038
26.5291 | -35.5885 | -36.8094
26.5352 | -35.5889 | -36.8151
26.5413 | -35.5892 | -36.8205
26.5474 | -35.5897 | -36.8261
26.5535 | -35.5901 | -36.8314
26.5596 | -35.5904 | -36.8367
26.5657 | -35.5909 | -36.8427
26.5718 | -35.5913 | -36.8486
26.578 | -35.5917 | -36.8549
26.5841 | -35.5921 | -36.8604
26.5902 | -35.5925 | -36.8657
26.5963 | -35.5929 | -36.8716
26.6024 | -35.5933 | -36.8774
26.6085 | -35.5937 | -36.8833
26.6146 | -35.5941 | -36.8883
26.6207 | -35.5945 | -36.8941
26.6268 | -35.5949 | -36.8997
26.6329 | -35.5954 | -36.9054
26.639 | -35.5958 | -36.9112
26.6451 | -35.5962 | -36.9169
26.6512 | -35.5965 | -36.922
26.6573 | -35.5969 | -36.928
26.6634 | -35.5974 | -36.9329
26.6695 | -35.5977 | -36.9388
26.6756 | -35.5982 | -36.9436
26.6817 | -35.5986 | -36.9489
26.6878 | -35.5989 | -36.9546
26.6939 | -35.5993 | -36.9602
26.7 | -35.5997 | -36.9658
26.7061 | -35.6001 | -36.9712
26.7122 | -35.6005 | -36.9771
26.7183 | -35.6008 | -36.9825
26.7245 | -35.6011 | -36.9878
26.7306 | -35.6015 | -36.9935
26.7367 | -35.6019 | -36.9993
26.7428 | -35.6023 | -37.0049
26.7489 | -35.6027 | -37.0105
26.755 | -35.6032 | -37.0162
26.7611 | -35.6035 | -37.0218
26.7672 | -35.6039 | -37.0273
26.7733 | -35.6044 | -37.033
26.7794 | -35.6048 | -37.0383
26.7855 | -35.6051 | -37.0438
26.7916 | -35.6055 | -37.0491
26.7977 | -35.6059 | -37.0545
26.8038 | -35.6063 | -37.0602
26.8099 | -35.6067 | -37.066
26.816 | -35.6071 | -37.0713
26.8221 | -35.6075 | -37.077
26.8282 | -35.6079 | -37.0826
26.8343 | -35.6084 | -37.0879
26.8404 | -35.6088 | -37.094
26.8465 | -35.6093 | -37.0994
26.8526 | -35.6096 | -37.104
26.8587 | -35.6101 | -37.1093
26.8649 | -35.6104 | -37.1149
26.871 | -35.6108 | -37.1207
26.8771 | -35.6112 | -37.1265
26.8832 | -35.6116 | -37.1318
26.8893 | -35.612 | -37.1379
26.8954 | -35.6124 | -37.1436
26.9015 | -35.6128 | -37.149
26.9076 | -35.6132 | -37.1545
26.9137 | -35.6136 | -37.1598
26.9198 | -35.614 | -37.1654
26.9259 | -35.6143 | -37.1709
26.932 | -35.6146 | -37.1763
26.9381 | -35.615 | -37.182
26.9442 | -35.6154 | -37.1872
26.9503 | -35.6158 | -37.1926
26.9564 | -35.6161 | -37.1984
26.9625 | -35.6165 | -37.204
26.9686 | -35.6169 | -37.2092
26.9747 | -35.6173 | -37.2147
26.9808 | -35.6178 | -37.2198
26.9869 | -35.6182 | -37.225
26.993 | -35.6187 | -37.2304
26.9991 | -35.6191 | -37.2355
27.0052 | -35.6196 | -37.2415
27.0114 | -35.62 | -37.2467
27.0175 | -35.6204 | -37.2516
27.0236 | -35.6208 | -37.2571
27.0297 | -35.6212 | -37.2624
27.0358 | -35.6216 | -37.268
27.0419 | -35.622 | -37.2736
27.048 | -35.6223 | -37.2787
27.0541 | -35.6227 | -37.2838
27.0602 | -35.6231 | -37.2894
27.0663 | -35.6235 | -37.2951
27.0724 | -35.624 | -37.3003
27.0785 | -35.6244 | -37.3057
27.0846 | -35.6248 | -37.3111
27.0907 | -35.6253 | -37.3166
27.0968 | -35.6257 | -37.3226
27.1029 | -35.626 | -37.328
27.109 | -35.6264 | -37.3336
27.1151 | -35.6268 | -37.3388
27.1212 | -35.6272 | -37.3442
27.1273 | -35.6277 | -37.3495
27.1334 | -35.6282 | -37.3548
27.1395 | -35.6287 | -37.3607
27.1456 | -35.6291 | -37.3661
27.1518 | -35.6295 | -37.3715
27.1579 | -35.6299 | -37.3771
27.164 | -35.6304 | -37.3821
27.1701 | -35.6307 | -37.3873
27.1762 | -35.6311 | -37.3928
27.1823 | -35.6315 | -37.3985
27.1884 | -35.6319 | -37.4043
27.1945 | -35.6323 | -37.4096
27.2006 | -35.6327 | -37.4153
27.2067 | -35.6331 | -37.4206
27.2128 | -35.6336 | -37.4263
27.2189 | -35.6339 | -37.4319
27.225 | -35.6344 | -37.4374
27.2311 | -35.6348 | -37.4421
27.2372 | -35.6352 | -37.4474
27.2433 | -35.6355 | -37.4523
27.2494 | -35.6359 | -37.4578
27.2555 | -35.6363 | -37.4632
27.2616 | -35.6367 | -37.4688
27.2677 | -35.6371 | -37.4744
27.2738 | -35.6375 | -37.4798
27.2799 | -35.6379 | -37.4857
27.286 | -35.6383 | -37.4908
27.2921 | -35.6387 | -37.4963
27.2983 | -35.6391 | -37.5018
27.3044 | -35.6396 | -37.507
27.3105 | -35.6399 | -37.5125
27.3166 | -35.6403 | -37.5183
27.3227 | -35.6407 | -37.5239
27.3288 | -35.6411 | -37.5295
27.3349 | -35.6414 | -37.5348
27.341 | -35.6418 | -37.54
27.3471 | -35.6422 | -37.5451
27.3532 | -35.6426 | -37.5503
27.3593 | -35.6429 | -37.5556
27.3654 | -35.6433 | -37.5612
27.3715 | -35.6438 | -37.5667
27.3776 | -35.6442 | -37.5722
27.3837 | -35.6445 | -37.5776
27.3898 | -35.645 | -37.5831
27.3959 | -35.6453 | -37.5884
27.402 | -35.6457 | -37.5937
27.4081 | -35.6462 | -37.5989
27.4142 | -35.6465 | -37.6047
27.4203 | -35.6469 | -37.6102
27.4264 | -35.6473 | -37.6157
27.4325 | -35.6478 | -37.6212
27.4387 | -35.6482 | -37.6269
27.4448 | -35.6486 | -37.632
27.4509 | -35.6489 | -37.6377
27.457 | -35.6493 | -37.6433
27.4631 | -35.6498 | -37.6492
27.4692 | -35.6501 | -37.6546
27.4753 | -35.6505 | -37.6594
27.4814 | -35.6509 | -37.665
27.4875 | -35.6513 | -37.6697
27.4936 | -35.6517 | -37.6752
27.4997 | -35.6521 | -37.6809
27.5058 | -35.6524 | -37.6863
27.5119 | -35.6529 | -37.692
27.518 | -35.6533 | -37.6973
27.5241 | -35.6537 | -37.7029
27.5302 | -35.6541 | -37.7085
27.5363 | -35.6545 | -37.7139
27.5424 | -35.655 | -37.7197
27.5485 | -35.6554 | -37.7251
27.5546 | -35.6558 | -37.7305
27.5607 | -35.6562 | -37.7361
27.5668 | -35.6566 | -37.7415
27.5729 | -35.657 | -37.7469
27.5791 | -35.6574 | -37.7525
27.5852 | -35.6577 | -37.7577
27.5913 | -35.6581 | -37.763
27.5974 | -35.6585 | -37.7683
27.6035 | -35.6589 | -37.774
27.6096 | -35.6592 | -37.7794
27.6157 | -35.6596 | -37.7846
27.6218 | -35.66 | -37.7901
27.6279 | -35.6604 | -37.7956
27.634 | -35.6608 | -37.8006
27.6401 | -35.6612 | -37.806
27.6462 | -35.6616 | -37.8114
27.6523 | -35.662 | -37.8168
27.6584 | -35.6624 | -37.8221
27.6645 | -35.6628 | -37.8273
27.6706 | -35.6632 | -37.8329
27.6767 | -35.6636 | -37.8385
27.6828 | -35.664 | -37.8439
27.6889 | -35.6644 | -37.8497
27.695 | -35.6647 | -37.8551
27.7011 | -35.6651 | -37.8604
27.7072 | -35.6654 | -37.8659
27.7133 | -35.6659 | -37.8712
27.7194 | -35.6663 | -37.8767
27.7256 | -35.6667 | -37.8821
27.7317 | -35.667 | -37.8874
27.7378 | -35.6674 | -37.8929
27.7439 | -35.6678 | -37.8985
27.75 | -35.6682 | -37.9041
27.7561 | -35.6685 | -37.9088
27.7622 | -35.6689 | -37.9141
27.7683 | -35.6693 | -37.9197
27.7744 | -35.6697 | -37.9251
27.7805 | -35.6701 | -37.9303
27.7866 | -35.6705 | -37.9358
27.7927 | -35.6708 | -37.9411
27.7988 | -35.6712 | -37.9461
27.8049 | -35.6716 | -37.9514
27.811 | -35.672 | -37.9573
27.8171 | -35.6724 | -37.9628
27.8232 | -35.6728 | -37.9683
27.8293 | -35.6732 | -37.9737
27.8354 | -35.6735 | -37.9794
27.8415 | -35.6739 | -37.9845
27.8476 | -35.6743 | -37.9898
27.8537 | -35.6747 | -37.9953
27.8598 | -35.6751 | -38.0004
27.866 | -35.6755 | -38.0063
27.8721 | -35.6759 | -38.0117
27.8782 | -35.6763 | -38.0172
27.8843 | -35.6767 | -38.0225
27.8904 | -35.6771 | -38.028
27.8965 | -35.6774 | -38.0332
27.9026 | -35.6778 | -38.0385
27.9087 | -35.6782 | -38.0439
27.9148 | -35.6785 | -38.0497
27.9209 | -35.6789 | -38.0552
27.927 | -35.6794 | -38.0603
27.9331 | -35.6797 | -38.0654
27.9392 | -35.6801 | -38.0706
27.9453 | -35.6804 | -38.0762
27.9514 | -35.6808 | -38.082
27.9575 | -35.6812 | -38.0873
27.9636 | -35.6816 | -38.0927
27.9697 | -35.6819 | -38.0981
27.9758 | -35.6823 | -38.1037
27.9819 | -35.6827 | -38.1091
27.988 | -35.6831 | -38.114
27.9941 | -35.6834 | -38.1193
28.0002 | -35.6837 | -38.1247
28.0063 | -35.6841 | -38.1303
28.0125 | -35.6845 | -38.1355
28.0186 | -35.6849 | -38.1409
28.0247 | -35.6852 | -38.1462
28.0308 | -35.6856 | -38.1519
28.0369 | -35.6861 | -38.1572
28.043 | -35.6864 | -38.1623
28.0491 | -35.6868 | -38.1673
28.0552 | -35.6872 | -38.1727
28.0613 | -35.6875 | -38.1781
28.0674 | -35.6878 | -38.1832
28.0735 | -35.6882 | -38.1892

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4351 | -42.8533
21.4908 | -35.5485 | -35.2797
16.4948 | -34.9238 | -27.8665
12.6602 | -34.5543 | -21.565
9.7171 | -34.3295 | -16.6319
7.4582 | -34.1788 | -12.6308
5.7244 | -34.0704 | -9.3857
4.3936 | -34.0088 | -6.7692
3.3722 | -33.9848 | -4.7119
2.5883 | -33.9852 | -3.148
1.9866 | -33.9851 | -1.8983
1.5248 | -33.9816 | -0.89563
1.1703 | -33.9963 | -0.24103
1 | -33.9912 | 0.084418
0.89822 | -33.9756 | 0.356
0.68941 | -33.9674 | 0.94037
0.52919 | -33.978 | 1.3934
0.40612 | -33.9781 | 2.0025
0.31171 | -33.9665 | 2.457
0.23928 | -33.9719 | 3.187
0.18358 | -33.9757 | 3.611
0.14091 | -33.9793 | 4.5047
0.10819 | -34.0049 | 5.6571
0.083017 | -34.0556 | 7.1296
0.063705 | -34.0793 | 9.0002
0.048923 | -34.1344 | 11.4897
0.037479 | -34.2596 | 14.6726
0.028801 | -34.447 | 18.677
0.022078 | -34.75 | 23.9911
0.017023 | -35.2344 | 30.3721
0.013018 | -36.0305 | 38.0345
0.010014 | -37.1496 | 45.9337

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003751 | -7.9021
0.007501 | -26.6136
0.011252 | -34.6955
0.015002 | -39.8853
0.018753 | -44.4901
0.022504 | -45.8608
0.026254 | -47.8777
0.030005 | -48.7909
0.033755 | -50.0508
0.037506 | -51.0102
0.041256 | -52.4022
0.045007 | -53.9419
0.048758 | -56.4533
0.052508 | -56.4016
0.056259 | -56.7494
0.060009 | -57.5693
0.06376 | -58.584
0.067511 | -58.6553
0.071261 | -59.3144
0.075012 | -60.7152
0.078762 | -60.3194
0.082513 | -60.363
0.086263 | -60.9832
0.090014 | -62.2413
0.093765 | -63.3037
0.097515 | -64.3467
0.10127 | -65.7201
0.10502 | -65.0068
0.10877 | -64.0549
0.11252 | -64.5492
0.11627 | -66.2116
0.12002 | -66.6651
0.12377 | -66.8042
0.12752 | -67.0665
0.13127 | -67.5514
0.13502 | -68.4161
0.13877 | -68.0947
0.14252 | -67.6689
0.14627 | -67.9854
0.15002 | -68.2068
0.15377 | -69.7087
0.15752 | -70.1956
0.16128 | -70.0125
0.16503 | -69.4965
0.16878 | -69.6765
0.17253 | -70.4888
0.17628 | -70.6778
0.18003 | -70.8525
0.18378 | -70.6833
0.18753 | -70.5193
0.19128 | -70.5964
0.19503 | -71.3362
0.19878 | -72.17
0.20253 | -72.7578
0.20628 | -72.4244
0.21003 | -72.6204
0.21378 | -72.6518
0.21753 | -72.5139
0.22129 | -73.1204
0.22504 | -72.8933
0.22879 | -73.1224
0.23254 | -72.4888
0.23629 | -72.1396
0.24004 | -72.6806
0.24379 | -72.7892
0.24754 | -73.3506
0.25129 | -74.2627
0.25504 | -75.8841
0.25879 | -75.3802
0.26254 | -74.4441
0.26629 | -74.3615
0.27004 | -74.5093
0.27379 | -74.1137
0.27754 | -74.7071
0.28129 | -75.2373
0.28504 | -75.157
0.2888 | -75.2176
0.29255 | -75.6914
0.2963 | -75.8616
0.30005 | -75.0878
0.3038 | -75.4864
0.30755 | -75.93
0.3113 | -76.387
0.31505 | -76.5679
0.3188 | -76.5281
0.32255 | -77.0246
0.3263 | -76.8874
0.33005 | -77.2102
0.3338 | -77.1961
0.33755 | -77.8725
0.3413 | -77.8355
0.34505 | -76.6064
0.3488 | -77.4937
0.35256 | -77.4814
0.35631 | -77.5713
0.36006 | -77.8299
0.36381 | -78.9802
0.36756 | -79.07
0.37131 | -79.122
0.37506 | -78.5254
0.37881 | -76.9891
0.38256 | -77.5892
0.38631 | -78.5588
0.39006 | -78.5974
0.39381 | -79.47
0.39756 | -78.8835
0.40131 | -78.1632
0.40506 | -77.9892
0.40881 | -78.1518
0.41256 | -79.1206
0.41631 | -79.4806
0.42007 | -78.9691
0.42382 | -79.2234
0.42757 | -79.0496
0.43132 | -78.9418
0.43507 | -80.0253
0.43882 | -80.0304
0.44257 | -80.3612
0.44632 | -80.6766
0.45007 | -79.9609
0.45382 | -79.1696
0.45757 | -81.2225
0.46132 | -81.0111
0.46507 | -79.1343
0.46882 | -78.9673
0.47257 | -79.9889
0.47632 | -79.3071
0.48007 | -79.4607
0.48383 | -80.6173
0.48758 | -81.1038
0.49133 | -81.8607
0.49508 | -82.2692
0.49883 | -81.1447
0.50258 | -80.6271
0.50633 | -80.3937
0.51008 | -80.5714
0.51383 | -80.8386
0.51758 | -81.597
0.52133 | -81.8387
0.52508 | -80.6168
0.52883 | -80.4177
0.53258 | -80.9864
0.53633 | -81.3472
0.54008 | -80.9276
0.54383 | -81.1309
0.54759 | -81.3535
0.55134 | -81.1159
0.55509 | -81.1977
0.55884 | -81.9114
0.56259 | -82.364
0.56634 | -82.0977
0.57009 | -81.7009
0.57384 | -81.5704
0.57759 | -81.8426
0.58134 | -82.4134
0.58509 | -82.6864
0.58884 | -82.6293
0.59259 | -82.7007
0.59634 | -82.9645
0.60009 | -83.1484
0.60384 | -82.9384
0.60759 | -82.354
0.61135 | -82.3993
0.6151 | -82.9129
0.61885 | -83.0918
0.6226 | -83.3751
0.62635 | -83.549
0.6301 | -83.0807
0.63385 | -82.4696
0.6376 | -82.2532
0.64135 | -82.1792
0.6451 | -82.2641
0.64885 | -82.9732
0.6526 | -83.7569
0.65635 | -84.238
0.6601 | -84.6401
0.66385 | -84.7008
0.6676 | -84.1956
0.67136 | -83.542
0.6751 | -83.2527
0.67886 | -83.2515
0.68261 | -83.2776
0.68636 | -83.5968
0.69011 | -84.117
0.69386 | -84.7543
0.69761 | -84.9388
0.70136 | -84.6012
0.70511 | -84.1933
0.70886 | -84.2021
0.71261 | -84.4543
0.71636 | -84.7883
0.72011 | -84.7731
0.72386 | -84.5651
0.72761 | -84.5782
0.73136 | -84.6254
0.73511 | -84.5267
0.73886 | -84.088
0.74262 | -83.9818
0.74637 | -84.3962
0.75012 | -84.7473
0.75387 | -85.1331
0.75762 | -85.417
0.76137 | -85.3156
0.76512 | -85.0638
0.76887 | -84.7342
0.77262 | -84.5562
0.77637 | -85.0648
0.78012 | -85.1138
0.78387 | -84.6523
0.78762 | -84.8566
0.79137 | -85.05
0.79512 | -84.9328
0.79888 | -85.089
0.80263 | -85.2063
0.80638 | -85.7408
0.81013 | -86.5768
0.81388 | -86.8535
0.81763 | -86.6312
0.82138 | -86.2446
0.82513 | -85.9988
0.82888 | -85.9315
0.83263 | -85.8229
0.83638 | -85.4614
0.84013 | -84.7203
0.84388 | -84.4733
0.84763 | -84.5804
0.85138 | -85.2343
0.85513 | -86.2542
0.85888 | -86.5009
0.86264 | -86.3703
0.86638 | -86.6327
0.87014 | -86.7203
0.87389 | -86.1653
0.87764 | -85.8704
0.88139 | -86.071
0.88514 | -86.3324
0.88889 | -86.1888
0.89264 | -85.7018
0.89639 | -85.6359
0.90014 | -86.2264
0.90389 | -86.7223
0.90764 | -86.5095
0.91139 | -85.9373
0.91514 | -85.7463
0.91889 | -86.0054
0.92264 | -86.459
0.92639 | -87.0755
0.93014 | -87.3966
0.9339 | -87.4377
0.93765 | -87.6196
0.9414 | -88.0961
0.94515 | -87.8186
0.9489 | -87.0632
0.95265 | -86.8199
0.9564 | -86.8954
0.96015 | -86.8727
0.9639 | -87.0044
0.96765 | -87.1742
0.9714 | -87.4186
0.97515 | -87.5116
0.9789 | -87.3327
0.98265 | -87.205
0.9864 | -86.8582
0.99016 | -86.6589
0.99391 | -86.4163
0.99766 | -85.5456
1.0014 | -85.2317
1.0052 | -85.9173
1.0089 | -86.577
1.0127 | -86.9405
1.0164 | -87.1846
1.0202 | -87.2284
1.0239 | -87.6037
1.0277 | -87.8707
1.0314 | -87.4478
1.0352 | -87.1879
1.0389 | -87.4164
1.0427 | -87.5065
1.0464 | -87.4949
1.0502 | -87.3602
1.0539 | -87.089
1.0577 | -87.1655
1.0614 | -87.6333
1.0652 | -87.9398
1.0689 | -87.8053
1.0727 | -87.4654
1.0764 | -87.6248
1.0802 | -88.2257
1.0839 | -88.4672
1.0877 | -88.215
1.0914 | -87.8861
1.0952 | -87.9331
1.0989 | -88.322
1.1027 | -88.704
1.1064 | -88.7355
1.1102 | -88.5445
1.1139 | -88.4819
1.1177 | -88.5976
1.1214 | -88.5642
1.1252 | -88.2352
1.1289 | -88.0137
1.1327 | -88.1118
1.1364 | -88.3494
1.1402 | -88.5389
1.1439 | -88.6245
1.1477 | -88.7208
1.1514 | -88.839
1.1552 | -88.845
1.1589 | -88.8089
1.1627 | -88.7707
1.1664 | -88.6843
1.1702 | -88.6854
1.1739 | -88.8078
1.1777 | -88.8423
1.1814 | -88.6833
1.1852 | -88.5532
1.1889 | -88.6472
1.1927 | -88.8243
1.1964 | -88.9216
1.2002 | -88.9997
1.2039 | -89.1149
1.2077 | -89.1919
1.2114 | -89.0612
1.2152 | -88.7105
1.2189 | -88.3954
1.2227 | -88.3715
1.2264 | -88.5908
1.2302 | -88.7645
1.2339 | -88.8095
1.2377 | -88.898
1.2414 | -89.0443
1.2452 | -89.174
1.2489 | -89.2597
1.2527 | -89.2633
1.2564 | -89.2134
1.2602 | -89.1948
1.2639 | -89.2746
1.2677 | -89.5156
1.2714 | -89.8722
1.2752 | -90.1668
1.279 | -90.2276
1.2827 | -90.0023
1.2865 | -89.6751
1.2902 | -89.5435
1.294 | -89.6881
1.2977 | -89.8709
1.3015 | -89.8782
1.3052 | -89.8379
1.309 | -89.8927
1.3127 | -90.013
1.3165 | -90.0828
1.3202 | -90.0894
1.324 | -90.1978
1.3277 | -90.4775
1.3315 | -90.7455
1.3352 | -90.7849
1.339 | -90.5676
1.3427 | -90.3569
1.3465 | -90.3928
1.3502 | -90.6015
1.354 | -90.5977
1.3577 | -90.1791
1.3615 | -89.7123
1.3652 | -89.6369
1.369 | -89.96
1.3727 | -90.2885
1.3765 | -90.2671
1.3802 | -90.0015
1.384 | -89.7949
1.3877 | -89.72
1.3915 | -89.739
1.3952 | -89.8355
1.399 | -89.9346
1.4027 | -89.9655
1.4065 | -90.0068
1.4102 | -90.2294
1.414 | -90.6813
1.4177 | -91.1689
1.4215 | -91.447
1.4252 | -91.4868
1.429 | -91.3988
1.4327 | -91.2738
1.4365 | -91.1521
1.4402 | -90.882
1.444 | -90.5052
1.4477 | -90.3865
1.4515 | -90.5889
1.4552 | -90.8759
1.459 | -91.0533
1.4627 | -90.9743
1.4665 | -90.6434
1.4702 | -90.295
1.474 | -90.1095
1.4777 | -90.1214
1.4815 | -90.2542
1.4852 | -90.3209
1.489 | -90.1904
1.4927 | -89.9536
1.4965 | -89.7569
1.5002 | -89.7771
1.504 | -90.1215
1.5077 | -90.529
1.5115 | -90.6668
1.5152 | -90.6933
1.519 | -90.834
1.5227 | -91.0544
1.5265 | -91.2471
1.5302 | -91.3667
1.534 | -91.4297
1.5377 | -91.3698
1.5415 | -91.0675
1.5452 | -90.6966
1.549 | -90.5718
1.5527 | -90.7223
1.5565 | -90.9105
1.5602 | -91.0251
1.564 | -91.1851
1.5677 | -91.4366
1.5715 | -91.6174
1.5752 | -91.4895
1.579 | -91.0718
1.5827 | -90.7993
1.5865 | -90.8923
1.5902 | -91.0371
1.594 | -90.9604
1.5978 | -90.8908
1.6015 | -91.0524
1.6053 | -91.2913
1.609 | -91.3368
1.6128 | -91.1625
1.6165 | -90.9707
1.6203 | -91.019
1.624 | -91.0804
1.6278 | -91.1868
1.6315 | -91.2693
1.6353 | -91.3331
1.639 | -91.3832
1.6428 | -91.4166
1.6465 | -91.4501
1.6503 | -91.5478
1.654 | -91.7117
1.6578 | -91.8713
1.6615 | -91.9452
1.6653 | -91.9797
1.669 | -92.049
1.6728 | -92.1933
1.6765 | -92.4086
1.6803 | -92.6423
1.684 | -92.7717
1.6878 | -92.7369
1.6915 | -92.6586
1.6953 | -92.6637
1.699 | -92.719
1.7028 | -92.6651
1.7065 | -92.4451
1.7103 | -92.1579
1.714 | -91.9795
1.7178 | -91.9638
1.7215 | -92.0393
1.7253 | -92.1381
1.729 | -92.2052
1.7328 | -92.2191
1.7365 | -92.2422
1.7403 | -92.3055
1.744 | -92.3594
1.7478 | -92.3375
1.7515 | -92.2887
1.7553 | -92.2921
1.759 | -92.3661
1.7628 | -92.4657
1.7665 | -92.5266
1.7703 | -92.5542
1.774 | -92.5976
1.7778 | -92.636
1.7815 | -92.5989
1.7853 | -92.4664
1.789 | -92.3198
1.7928 | -92.2765
1.7965 | -92.4141
1.8003 | -92.6568
1.804 | -92.8014
1.8078 | -92.6896
1.8115 | -92.3654
1.8153 | -92.0324
1.819 | -91.8564
1.8228 | -91.8276
1.8265 | -91.8633
1.8303 | -91.9053
1.834 | -91.9804
1.8378 | -92.1155
1.8415 | -92.2955
1.8453 | -92.4794
1.849 | -92.5935
1.8528 | -92.5975
1.8565 | -92.5194
1.8603 | -92.467
1.864 | -92.5413
1.8678 | -92.6944
1.8715 | -92.8067
1.8753 | -92.8678
1.879 | -92.9846
1.8828 | -93.2155
1.8865 | -93.4818
1.8903 | -93.6532
1.894 | -93.6294
1.8978 | -93.4195
1.9015 | -93.1486
1.9053 | -92.9295
1.909 | -92.7742
1.9128 | -92.6365
1.9165 | -92.5582
1.9203 | -92.6505
1.9241 | -92.9805
1.9278 | -93.4334
1.9316 | -93.7608
1.9353 | -93.7936
1.9391 | -93.6113
1.9428 | -93.3715
1.9466 | -93.1485
1.9503 | -92.8995
1.9541 | -92.6325
1.9578 | -92.4727
1.9616 | -92.5523
1.9653 | -92.8789
1.9691 | -93.3238
1.9728 | -93.6602
1.9766 | -93.7542
1.9803 | -93.6299
1.9841 | -93.4152
1.9878 | -93.2125
1.9916 | -93.061
1.9953 | -92.9943
1.9991 | -93.0177
2.0028 | -93.1207
2.0066 | -93.2747
2.0103 | -93.3897
2.0141 | -93.3851
2.0178 | -93.2407
2.0216 | -93.0315
2.0253 | -92.894
2.0291 | -92.8958
2.0328 | -93.0246
2.0366 | -93.2425
2.0403 | -93.4949
2.0441 | -93.7057
2.0478 | -93.815
2.0516 | -93.8047
2.0553 | -93.7141
2.0591 | -93.6153
2.0628 | -93.5793
2.0666 | -93.6209
2.0703 | -93.7292
2.0741 | -93.8591
2.0778 | -93.9508
2.0816 | -93.9392
2.0853 | -93.8501
2.0891 | -93.7678
2.0928 | -93.7565
2.0966 | -93.8001
2.1003 | -93.835
2.1041 | -93.8043
2.1078 | -93.7145
2.1116 | -93.554
2.1153 | -93.3518
2.1191 | -93.1511
2.1228 | -93.0438
2.1266 | -93.113
2.1303 | -93.3457
2.1341 | -93.6558
2.1378 | -93.9483
2.1416 | -94.1724
2.1453 | -94.334
2.1491 | -94.4315
2.1528 | -94.4599
2.1566 | -94.395
2.1603 | -94.2784
2.1641 | -94.1539
2.1678 | -94.0228
2.1716 | -93.8591
2.1753 | -93.6836
2.1791 | -93.5412
2.1828 | -93.484
2.1866 | -93.5387
2.1903 | -93.7087
2.1941 | -93.9844
2.1978 | -94.3217
2.2016 | -94.6265
2.2053 | -94.8067
2.2091 | -94.833
2.2128 | -94.7342
2.2166 | -94.5678
2.2203 | -94.4067
2.2241 | -94.3103
2.2278 | -94.3013
2.2316 | -94.3576
2.2353 | -94.461
2.2391 | -94.5802
2.2428 | -94.7083
2.2466 | -94.8257
2.2504 | -94.919
2.2541 | -94.9712
2.2579 | -94.9938
2.2616 | -94.9555
2.2654 | -94.8314
2.2691 | -94.6445
2.2729 | -94.4513
2.2766 | -94.3042
2.2804 | -94.2489
2.2841 | -94.2697
2.2879 | -94.3333
2.2916 | -94.3997
2.2954 | -94.4036
2.2991 | -94.2834
2.3029 | -94.1039
2.3066 | -93.9446
2.3104 | -93.8558
2.3141 | -93.8513
2.3179 | -93.8933
2.3216 | -93.9188
2.3254 | -93.9071
2.3291 | -93.8531
2.3329 | -93.8014
2.3366 | -93.8063
2.3404 | -93.8986
2.3441 | -94.0696
2.3479 | -94.2972
2.3516 | -94.5436
2.3554 | -94.7257
2.3591 | -94.7766
2.3629 | -94.7103
2.3666 | -94.5631
2.3704 | -94.398
2.3741 | -94.2446
2.3779 | -94.1306
2.3816 | -94.0537
2.3854 | -94.0265
2.3891 | -94.0554
2.3929 | -94.1301
2.3966 | -94.2424
2.4004 | -94.3784
2.4041 | -94.5068
2.4079 | -94.6036
2.4116 | -94.6379
2.4154 | -94.6209
2.4191 | -94.5833
2.4229 | -94.57
2.4266 | -94.5904
2.4304 | -94.6179
2.4341 | -94.622
2.4379 | -94.5916
2.4416 | -94.5351
2.4454 | -94.485
2.4491 | -94.4718
2.4529 | -94.4859
2.4566 | -94.4775
2.4604 | -94.4387
2.4641 | -94.3726
2.4679 | -94.3043
2.4716 | -94.2959
2.4754 | -94.3727
2.4791 | -94.5267
2.4829 | -94.7223
2.4866 | -94.886
2.4904 | -94.9497
2.4941 | -94.8977
2.4979 | -94.8067
2.5016 | -94.7675
2.5054 | -94.8348
2.5091 | -94.9799
2.5129 | -95.1065
2.5166 | -95.1347
2.5204 | -95.0462
2.5241 | -94.9217
2.5279 | -94.8687
2.5316 | -94.9383
2.5354 | -95.1135
2.5391 | -95.3225
2.5429 | -95.5106
2.5466 | -95.6555
2.5504 | -95.7483
2.5541 | -95.7818
2.5579 | -95.749
2.5617 | -95.6219
2.5654 | -95.4019
2.5692 | -95.142
2.5729 | -94.903
2.5767 | -94.7315
2.5804 | -94.6479
2.5842 | -94.6479
2.5879 | -94.7144
2.5917 | -94.8428
2.5954 | -95.0295
2.5992 | -95.2641
2.6029 | -95.5066
2.6067 | -95.6991
2.6104 | -95.8092
2.6142 | -95.88
2.6179 | -95.9657
2.6217 | -96.0736
2.6254 | -96.1675
2.6292 | -96.2069
2.6329 | -96.1738
2.6367 | -96.0694
2.6404 | -95.8899
2.6442 | -95.6601
2.6479 | -95.4226
2.6517 | -95.2357
2.6554 | -95.1143
2.6592 | -95.0482
2.6629 | -95.0375
2.6667 | -95.0953
2.6704 | -95.2077
2.6742 | -95.3198
2.6779 | -95.3932
2.6817 | -95.4148
2.6854 | -95.4056
2.6892 | -95.4193
2.6929 | -95.4917
2.6967 | -95.6035
2.7004 | -95.6613
2.7042 | -95.4897
2.7079 | -95.3787
2.7117 | -95.2667
2.7154 | -95.2274
2.7192 | -95.2985
2.7229 | -95.4428
2.7267 | -95.6131
2.7304 | -95.7677
2.7342 | -95.8775
2.7379 | -95.9563
2.7417 | -96.024
2.7454 | -96.082
2.7492 | -96.1098
2.7529 | -96.091
2.7567 | -96.0076
2.7604 | -95.8915
2.7642 | -95.7722
2.7679 | -95.6857
2.7717 | -95.6606
2.7754 | -95.6808
2.7792 | -95.7173
2.7829 | -95.7253
2.7867 | -95.6737
2.7904 | -95.5744
2.7942 | -95.4742
2.7979 | -95.4153
2.8017 | -95.4264
2.8054 | -95.5169
2.8092 | -95.6451
2.8129 | -95.7921
2.8167 | -95.9275
2.8204 | -96.0354
2.8242 | -96.1089
2.8279 | -96.1455
2.8317 | -96.1465
2.8354 | -96.1168
2.8392 | -96.0844
2.8429 | -96.0816
2.8467 | -96.1356
2.8504 | -96.2404
2.8542 | -96.3524
2.8579 | -96.4331
2.8617 | -96.4375
2.8654 | -96.3875
2.8692 | -96.3135
2.8729 | -96.2512
2.8767 | -96.2155
2.8805 | -96.2198
2.8842 | -96.2583
2.888 | -96.3144
2.8917 | -96.3892
2.8955 | -96.4747
2.8992 | -96.5573
2.903 | -96.6224
2.9067 | -96.6541
2.9105 | -96.6606
2.9142 | -96.631
2.918 | -96.5792
2.9217 | -96.5325
2.9255 | -96.5186
2.9292 | -96.5218
2.933 | -96.5312
2.9367 | -96.5541
2.9405 | -96.5763
2.9442 | -96.5873
2.948 | -96.5902
2.9517 | -96.5682
2.9555 | -96.5157
2.9592 | -96.4321
2.963 | -96.3595
2.9667 | -96.33
2.9705 | -96.365
2.9742 | -96.4526
2.978 | -96.5745
2.9817 | -96.6922
2.9855 | -96.7415
2.9892 | -96.7137
2.993 | -96.6342
2.9967 | -96.5525
3.0005 | -96.5189
3.0042 | -96.5429
3.008 | -96.5984
3.0117 | -96.6451
3.0155 | -96.6407
3.0192 | -96.5786
3.023 | -96.4926
3.0267 | -96.4208
3.0305 | -96.3868
3.0342 | -96.4182
3.038 | -96.4947
3.0417 | -96.5874
3.0455 | -96.664
3.0492 | -96.7273
3.053 | -96.7921
3.0567 | -96.8755
3.0605 | -96.9714
3.0642 | -97.0679
3.068 | -97.1321
3.0717 | -97.1031
3.0755 | -96.9894
3.0792 | -96.8274
3.083 | -96.671
3.0867 | -96.5654
3.0905 | -96.5326
3.0942 | -96.5827
3.098 | -96.6934
3.1017 | -96.8201
3.1055 | -96.9064
3.1092 | -96.9143
3.113 | -96.8327
3.1167 | -96.6918
3.1205 | -96.5526
3.1242 | -96.4318
3.128 | -96.3636
3.1317 | -96.344
3.1355 | -96.3483
3.1392 | -96.3417
3.143 | -96.2992
3.1467 | -96.2412
3.1505 | -96.2156
3.1542 | -96.2585
3.158 | -96.3881
3.1617 | -96.6024
3.1655 | -96.8755
3.1692 | -97.1689
3.173 | -97.4417
3.1767 | -97.6538
3.1805 | -97.779
3.1842 | -97.8156
3.188 | -97.7467
3.1917 | -97.5931
3.1955 | -97.3879
3.1993 | -97.1549
3.203 | -96.9246
3.2068 | -96.7274
3.2105 | -96.5782
3.2143 | -96.4935
3.218 | -96.4683
3.2218 | -96.4718
3.2255 | -96.4861
3.2293 | -96.513
3.233 | -96.5637
3.2368 | -96.6703
3.2405 | -96.8289
3.2443 | -96.9768
3.248 | -97.0923
3.2518 | -97.1737
3.2555 | -97.22
3.2593 | -97.2467
3.263 | -97.2772
3.2668 | -97.3273
3.2705 | -97.404
3.2743 | -97.4997
3.278 | -97.5785
3.2818 | -97.6302
3.2855 | -97.647
3.2893 | -97.6256
3.293 | -97.5748
3.2968 | -97.5181
3.3005 | -97.4685
3.3043 | -97.4307
3.308 | -97.399
3.3118 | -97.3663
3.3155 | -97.3226
3.3193 | -97.2835
3.323 | -97.2512
3.3268 | -97.2312
3.3305 | -97.2333
3.3343 | -97.244
3.338 | -97.2508
3.3418 | -97.2528
3.3455 | -97.2489
3.3493 | -97.2267
3.353 | -97.1782
3.3568 | -97.1004
3.3605 | -97.0002
3.3643 | -96.892
3.368 | -96.8013
3.3718 | -96.7513
3.3755 | -96.7561
3.3793 | -96.8229
3.383 | -96.9487
3.3868 | -97.1028
3.3905 | -97.2548
3.3943 | -97.3691
3.398 | -97.4163
3.4018 | -97.4007
3.4055 | -97.3483
3.4093 | -97.2856
3.413 | -97.2415
3.4168 | -97.2438
3.4205 | -97.2916
3.4243 | -97.3661
3.428 | -97.4466
3.4318 | -97.5116
3.4355 | -97.5427
3.4393 | -97.529
3.443 | -97.4719
3.4468 | -97.3907
3.4505 | -97.3043
3.4543 | -97.2401
3.458 | -97.21
3.4618 | -97.2079
3.4655 | -97.2133
3.4693 | -97.2106
3.473 | -97.1813
3.4768 | -97.1407
3.4805 | -97.1101
3.4843 | -97.0989
3.488 | -97.1112
3.4918 | -97.1592
3.4955 | -97.2289
3.4993 | -97.3118
3.503 | -97.3924
3.5068 | -97.4585
3.5105 | -97.507
3.5143 | -97.5502
3.5181 | -97.6078
3.5218 | -97.6885
3.5256 | -97.7846
3.5293 | -97.8855
3.5331 | -97.9681
3.5368 | -98.017
3.5406 | -98.0374
3.5443 | -98.0184
3.5481 | -97.9597
3.5518 | -97.8822
3.5556 | -97.7735
3.5593 | -97.6391
3.5631 | -97.5039
3.5668 | -97.3812
3.5706 | -97.2807
3.5743 | -97.2261
3.5781 | -97.2322
3.5818 | -97.308
3.5856 | -97.4524
3.5893 | -97.643
3.5931 | -97.8456
3.5968 | -98.0269
3.6006 | -98.1759
3.6043 | -98.2759
3.6081 | -98.3308
3.6118 | -98.3747
3.6156 | -98.4083
3.6193 | -98.4307
3.6231 | -98.4466
3.6268 | -98.4441
3.6306 | -98.4075
3.6343 | -98.3288
3.6381 | -98.2206
3.6418 | -98.0886
3.6456 | -97.9507
3.6493 | -97.8069
3.6531 | -97.6707
3.6568 | -97.554
3.6606 | -97.4754
3.6643 | -97.4413
3.6681 | -97.4616
3.6718 | -97.5384
3.6756 | -97.6534
3.6793 | -97.7914
3.6831 | -97.9236
3.6868 | -98.0222
3.6906 | -98.0845
3.6943 | -98.1025
3.6981 | -98.1066
3.7018 | -98.1206
3.7056 | -98.1718
3.7093 | -98.2501
3.7131 | -98.3495
3.7168 | -98.4478
3.7206 | -98.5312
3.7243 | -98.5911
3.7281 | -98.6308
3.7318 | -98.6556
3.7356 | -98.6677
3.7393 | -98.6653
3.7431 | -98.6431
3.7468 | -98.5921
3.7506 | -98.5127
3.7543 | -98.4036
3.7581 | -98.2883
3.7618 | -98.1591
3.7656 | -98.0356
3.7693 | -97.9292
3.7731 | -97.8469
3.7768 | -97.7793
3.7806 | -97.7196
3.7843 | -97.7438
3.7881 | -97.7678
3.7918 | -97.8396
3.7956 | -97.9579
3.7993 | -98.1097
3.8031 | -98.2815
3.8068 | -98.4522
3.8106 | -98.5895
3.8143 | -98.675
3.8181 | -98.6967
3.8218 | -98.672
3.8256 | -98.6082
3.8293 | -98.5226
3.8331 | -98.4395
3.8369 | -98.3901
3.8406 | -98.3785
3.8444 | -98.4022
3.8481 | -98.4521
3.8519 | -98.5024
3.8556 | -98.5287
3.8594 | -98.5267
3.8631 | -98.4999
3.8669 | -98.4521
3.8706 | -98.3965
3.8744 | -98.356
3.8781 | -98.3316
3.8819 | -98.3123
3.8856 | -98.304
3.8894 | -98.3053
3.8931 | -98.3007
3.8969 | -98.2942
3.9006 | -98.284
3.9044 | -98.2719
3.9081 | -98.2561
3.9119 | -98.2431
3.9156 | -98.2465
3.9194 | -98.2586
3.9231 | -98.2701
3.9269 | -98.282
3.9306 | -98.2938
3.9344 | -98.3159
3.9381 | -98.3513
3.9419 | -98.3916
3.9456 | -98.4513
3.9494 | -98.5055
3.9531 | -98.5405
3.9569 | -98.5754
3.9606 | -98.5876
3.9644 | -98.5733
3.9681 | -98.5419
3.9719 | -98.5069
3.9756 | -98.4722
3.9794 | -98.4327
3.9831 | -98.3954
3.9869 | -98.3734
3.9906 | -98.3529
3.9944 | -98.3376
3.9981 | -98.3254
4.0019 | -98.311
4.0056 | -98.3025
4.0094 | -98.2932
4.0131 | -98.2955
4.0169 | -98.2991
4.0206 | -98.293
4.0244 | -98.2986
4.0281 | -98.3094
4.0319 | -98.3186
4.0356 | -98.3311
4.0394 | -98.3479
4.0431 | -98.3696
4.0469 | -98.3986
4.0506 | -98.4318
4.0544 | -98.4817
4.0581 | -98.541
4.0619 | -98.6177
4.0656 | -98.7035
4.0694 | -98.7866
4.0731 | -98.8553
4.0769 | -98.9056
4.0806 | -98.9491
4.0844 | -98.9972
4.0881 | -99.0473
4.0919 | -99.114
4.0956 | -99.1863
4.0994 | -99.2279
4.1031 | -99.2377
4.1069 | -99.2168
4.1106 | -99.1675
4.1144 | -99.111
4.1181 | -99.0591
4.1219 | -99.0099
4.1256 | -98.9767
4.1294 | -98.9675
4.1331 | -98.9666
4.1369 | -98.9708
4.1406 | -98.9775
4.1444 | -98.9797
4.1481 | -98.9765
4.1519 | -98.981
4.1556 | -98.9918
4.1594 | -99.0157
4.1631 | -99.0562
4.1669 | -99.0991
4.1707 | -99.1371
4.1744 | -99.1633
4.1782 | -99.1697
4.1819 | -99.1695
4.1857 | -99.1612
4.1894 | -99.1319
4.1932 | -99.0937
4.1969 | -99.0606
4.2007 | -99.0286
4.2044 | -99.0022
4.2082 | -98.9804
4.2119 | -98.9604
4.2157 | -98.9485
4.2194 | -98.9365
4.2232 | -98.9363
4.2269 | -98.9532
4.2307 | -98.9818
4.2344 | -99.0108
4.2382 | -99.0359
4.2419 | -99.0502
4.2457 | -99.0536
4.2494 | -99.0607
4.2532 | -99.0806
4.2569 | -99.1142
4.2607 | -99.1656
4.2644 | -99.2328
4.2682 | -99.3004
4.2719 | -99.3588
4.2757 | -99.3946
4.2794 | -99.419
4.2832 | -99.4464
4.2869 | -99.4593
4.2907 | -99.4704
4.2944 | -99.4864
4.2982 | -99.51
4.3019 | -99.5373
4.3057 | -99.5762
4.3094 | -99.6206
4.3132 | -99.6675
4.3169 | -99.7205
4.3207 | -99.7723
4.3244 | -99.8078
4.3282 | -99.809
4.3319 | -99.8096
4.3357 | -99.7998
4.3394 | -99.7744
4.3432 | -99.7319
4.3469 | -99.6652
4.3507 | -99.582
4.3544 | -99.4885
4.3582 | -99.3858
4.3619 | -99.2888
4.3657 | -99.2011
4.3694 | -99.1436
4.3732 | -99.1176
4.3769 | -99.124
4.3807 | -99.156
4.3844 | -99.2043
4.3882 | -99.2627
4.3919 | -99.3249
4.3957 | -99.3846
4.3994 | -99.449
4.4032 | -99.5084
4.4069 | -99.5661
4.4107 | -99.6242
4.4144 | -99.6906
4.4182 | -99.7543
4.4219 | -99.819
4.4257 | -99.8838
4.4294 | -99.9338
4.4332 | -99.9726
4.4369 | -100.0049
4.4407 | -100.0114
4.4444 | -100.0079
4.4482 | -99.9951
4.4519 | -99.9735
4.4557 | -99.9459
4.4594 | -99.9078
4.4632 | -99.8696
4.4669 | -99.8343
4.4707 | -99.8057
4.4744 | -99.7899
4.4782 | -99.7888
4.482 | -99.7986
4.4857 | -99.8092
4.4895 | -99.8254
4.4932 | -99.8357
4.497 | -99.8405
4.5007 | -99.8402
4.5045 | -99.8224
4.5082 | -99.7921
4.512 | -99.7548
4.5157 | -99.7079
4.5195 | -99.6703
4.5232 | -99.6543
4.527 | -99.6575
4.5307 | -99.691
4.5345 | -99.7423
4.5382 | -99.8089
4.542 | -99.881
4.5457 | -99.9505
4.5495 | -100.012
4.5532 | -100.0559
4.557 | -100.0845
4.5607 | -100.0902
4.5645 | -100.0765
4.5682 | -100.0389
4.572 | -99.9905
4.5757 | -99.9386
4.5795 | -99.8841
4.5832 | -99.8386
4.587 | -99.8059
4.5907 | -99.7872
4.5945 | -99.7879
4.5982 | -99.8079
4.602 | -99.8433
4.6057 | -99.8942
4.6095 | -99.9511
4.6132 | -100.0148
4.617 | -100.0767
4.6207 | -100.1255
4.6245 | -100.157
4.6282 | -100.17
4.632 | -100.1783
4.6357 | -100.1733
4.6395 | -100.1614
4.6432 | -100.1437
4.647 | -100.121
4.6507 | -100.0908
4.6545 | -100.0534
4.6582 | -100.0031
4.662 | -99.9497
4.6657 | -99.8907
4.6695 | -99.8415
4.6732 | -99.8153
4.677 | -99.8091
4.6807 | -99.8204
4.6845 | -99.8415
4.6882 | -99.8794
4.692 | -99.9349
4.6957 | -100.0002
4.6995 | -100.0644
4.7032 | -100.1261
4.707 | -100.1781
4.7107 | -100.2161
4.7145 | -100.2394
4.7182 | -100.236
4.722 | -100.2188
4.7257 | -100.1766
4.7295 | -100.1221
4.7332 | -100.0588
4.737 | -99.9969
4.7407 | -99.9416
4.7445 | -99.8957
4.7482 | -99.8688
4.752 | -99.8595
4.7557 | -99.8735
4.7595 | -99.9016
4.7632 | -99.9373
4.767 | -99.9826
4.7707 | -100.0269
4.7745 | -100.0582
4.7782 | -100.0775
4.782 | -100.076
4.7857 | -100.0464
4.7895 | -100.0199
4.7932 | -99.9935
4.797 | -99.9815
4.8007 | -99.9781
4.8045 | -99.9886
4.8083 | -100.0134
4.812 | -100.0433
4.8158 | -100.0813
4.8195 | -100.1235
4.8233 | -100.1759
4.827 | -100.2265
4.8308 | -100.2783
4.8345 | -100.3367
4.8383 | -100.3907
4.842 | -100.4409
4.8458 | -100.4931
4.8495 | -100.5279
4.8533 | -100.5388
4.857 | -100.5356
4.8608 | -100.4947
4.8645 | -100.4249
4.8683 | -100.3139
4.872 | -100.2289
4.8758 | -100.1509
4.8795 | -100.0744
4.8833 | -100.0107
4.887 | -99.9755
4.8908 | -99.9672
4.8945 | -99.9849
4.8983 | -100.0254
4.902 | -100.0944
4.9058 | -100.1868
4.9095 | -100.2892
4.9133 | -100.3937
4.917 | -100.486
4.9208 | -100.5517
4.9245 | -100.5936
4.9283 | -100.6107
4.932 | -100.5995
4.9358 | -100.5767
4.9395 | -100.5451
4.9433 | -100.5052
4.947 | -100.4747
4.9508 | -100.4491
4.9545 | -100.4311
4.9583 | -100.4306
4.962 | -100.4498
4.9658 | -100.4676
4.9695 | -100.4862
4.9733 | -100.5063
4.977 | -100.5093
4.9808 | -100.4849
4.9845 | -100.4381
4.9883 | -100.371
4.992 | -100.3006
4.9958 | -100.2398
4.9995 | -100.1917
5.0033 | -100.1552
5.007 | -100.1227
5.0108 | -100.0995
5.0145 | -100.0777
5.0183 | -100.0609
5.022 | -100.0548
5.0258 | -100.0605
5.0295 | -100.0874
5.0333 | -100.1294
5.037 | -100.1877
5.0408 | -100.2566
5.0445 | -100.3442
5.0483 | -100.4388
5.052 | -100.5258
5.0558 | -100.6063
5.0595 | -100.6737
5.0633 | -100.7224
5.067 | -100.7575
5.0708 | -100.7747
5.0745 | -100.7796
5.0783 | -100.7836
5.082 | -100.7996
5.0858 | -100.8336
5.0895 | -100.876
5.0933 | -100.9268
5.097 | -100.9743
5.1008 | -101.0114
5.1045 | -101.0285
5.1083 | -101.0297
5.112 | -101.025
5.1158 | -101.0053
5.1196 | -100.9749
5.1233 | -100.9445
5.1271 | -100.9183
5.1308 | -100.8997
5.1346 | -100.8807
5.1383 | -100.8676
5.1421 | -100.8543
5.1458 | -100.839
5.1496 | -100.8319
5.1533 | -100.8392
5.1571 | -100.8478
5.1608 | -100.8569
5.1646 | -100.8615
5.1683 | -100.8588
5.1721 | -100.8493
5.1758 | -100.835
5.1796 | -100.8207
5.1833 | -100.8152
5.1871 | -100.8144
5.1908 | -100.824
5.1946 | -100.8456
5.1983 | -100.879
5.2021 | -100.9193
5.2058 | -100.9666
5.2096 | -101.0288
5.2133 | -101.1011
5.2171 | -101.1646
5.2208 | -101.2276
5.2246 | -101.2815
5.2283 | -101.3026
5.2321 | -101.2919
5.2358 | -101.2613
5.2396 | -101.2032
5.2433 | -101.1233
5.2471 | -101.0346
5.2508 | -100.9454
5.2546 | -100.8535
5.2583 | -100.7671
5.2621 | -100.6895
5.2658 | -100.6207
5.2696 | -100.5613
5.2733 | -100.5132
5.2771 | -100.4779
5.2808 | -100.4592
5.2846 | -100.4549
5.2883 | -100.4631
5.2921 | -100.4856
5.2958 | -100.5243
5.2996 | -100.5726
5.3033 | -100.6252
5.3071 | -100.6853
5.3108 | -100.7489
5.3146 | -100.8229
5.3183 | -100.899
5.3221 | -100.9739
5.3258 | -101.0389
5.3296 | -101.0871
5.3333 | -101.1287
5.3371 | -101.1523
5.3408 | -101.1588
5.3446 | -101.1638
5.3483 | -101.1655
5.3521 | -101.1731
5.3558 | -101.1889
5.3596 | -101.2069
5.3633 | -101.2319
5.3671 | -101.2543
5.3708 | -101.2669
5.3746 | -101.2647
5.3783 | -101.243
5.3821 | -101.2052
5.3858 | -101.1539
5.3896 | -101.0838
5.3933 | -101.0055
5.3971 | -100.9278
5.4008 | -100.8579
5.4046 | -100.793
5.4083 | -100.775
5.4121 | -100.7467
5.4158 | -100.7267
5.4196 | -100.7209
5.4233 | -100.7304
5.4271 | -100.7525
5.4308 | -100.7757
5.4346 | -100.7949
5.4383 | -100.8185
5.4421 | -100.8423
5.4459 | -100.8627
5.4496 | -100.8819
5.4534 | -100.9024
5.4571 | -100.9267
5.4609 | -100.9478
5.4646 | -100.9725
5.4684 | -101.0008
5.4721 | -101.0429
5.4759 | -101.0873
5.4796 | -101.1392
5.4834 | -101.1891
5.4871 | -101.2444
5.4909 | -101.2897
5.4946 | -101.3215
5.4984 | -101.3432
5.5021 | -101.3562
5.5059 | -101.353
5.5096 | -101.328
5.5134 | -101.2872
5.5171 | -101.25
5.5209 | -101.212
5.5246 | -101.1668
5.5284 | -101.1318
5.5321 | -101.1061
5.5359 | -101.0933
5.5396 | -101.0888
5.5434 | -101.0972
5.5471 | -101.1251
5.5509 | -101.1571
5.5546 | -101.1852
5.5584 | -101.2069
5.5621 | -101.2188
5.5659 | -101.2231
5.5696 | -101.2235
5.5734 | -101.2136
5.5771 | -101.2002
5.5809 | -101.1831
5.5846 | -101.1679
5.5884 | -101.1595
5.5921 | -101.1617
5.5959 | -101.1768
5.5996 | -101.1971
5.6034 | -101.2236
5.6071 | -101.2599
5.6109 | -101.3098
5.6146 | -101.3607
5.6184 | -101.4145
5.6221 | -101.465
5.6259 | -101.5158
5.6296 | -101.5614
5.6334 | -101.6041
5.6371 | -101.6452
5.6409 | -101.6833
5.6446 | -101.7084
5.6484 | -101.7235
5.6521 | -101.7259
5.6559 | -101.7182
5.6596 | -101.7138
5.6634 | -101.7083
5.6671 | -101.7062
5.6709 | -101.7046
5.6746 | -101.7006
5.6784 | -101.6978
5.6821 | -101.6973
5.6859 | -101.6994
5.6896 | -101.7027
5.6934 | -101.6993
5.6971 | -101.7002
5.7009 | -101.7089
5.7046 | -101.7191
5.7084 | -101.737
5.7121 | -101.7548
5.7159 | -101.7676
5.7196 | -101.7758
5.7234 | -101.7777
5.7271 | -101.7752
5.7309 | -101.769
5.7346 | -101.7519
5.7384 | -101.7217
5.7421 | -101.6864
5.7459 | -101.6452
5.7496 | -101.6066
5.7534 | -101.5809
5.7572 | -101.5679
5.7609 | -101.5694
5.7647 | -101.5875
5.7684 | -101.6124
5.7722 | -101.6432
5.7759 | -101.6766
5.7797 | -101.7134
5.7834 | -101.7404
5.7872 | -101.757
5.7909 | -101.7568
5.7947 | -101.7425
5.7984 | -101.7179
5.8022 | -101.6816
5.8059 | -101.6392
5.8097 | -101.5905
5.8134 | -101.5395
5.8172 | -101.4953
5.8209 | -101.4683
5.8247 | -101.4513
5.8284 | -101.4344
5.8322 | -101.4284
5.8359 | -101.4374
5.8397 | -101.4481
5.8434 | -101.4573
5.8472 | -101.4707
5.8509 | -101.4818
5.8547 | -101.4994
5.8584 | -101.5206
5.8622 | -101.5394
5.8659 | -101.5625
5.8697 | -101.5788
5.8734 | -101.5919
5.8772 | -101.5934
5.8809 | -101.5718
5.8847 | -101.5277
5.8884 | -101.4695
5.8922 | -101.4003
5.8959 | -101.3248
5.8997 | -101.241
5.9034 | -101.1598
5.9072 | -101.0868
5.9109 | -101.0345
5.9147 | -101.0099
5.9184 | -101.0123
5.9222 | -101.0377
5.9259 | -101.0943
5.9297 | -101.1827
5.9334 | -101.2867
5.9372 | -101.3988
5.9409 | -101.5067
5.9447 | -101.6073
5.9484 | -101.6548
5.9522 | -101.7166
5.9559 | -101.7548
5.9597 | -101.7792
5.9634 | -101.787
5.9672 | -101.7877
5.9709 | -101.7849
5.9747 | -101.7784
5.9784 | -101.7645
5.9822 | -101.7415
5.9859 | -101.7231
5.9897 | -101.7161
5.9934 | -101.7089
5.9972 | -101.7031
6.0009 | -101.7028
6.0047 | -101.711
6.0084 | -101.7232
6.0122 | -101.7435
6.0159 | -101.7745
6.0197 | -101.8086
6.0234 | -101.8482
6.0272 | -101.8932
6.0309 | -101.9385
6.0347 | -101.9817
6.0384 | -102.0199
6.0422 | -102.0552
6.0459 | -102.0842
6.0497 | -102.1001
6.0534 | -102.0989
6.0572 | -102.0948
6.0609 | -102.084
6.0647 | -102.0698
6.0684 | -102.0551
6.0722 | -102.0449
6.0759 | -102.0432
6.0797 | -102.0383
6.0835 | -102.0279
6.0872 | -102.0193
6.091 | -102.006
6.0947 | -101.987
6.0985 | -101.9618
6.1022 | -101.936
6.106 | -101.9151
6.1097 | -101.8906
6.1135 | -101.8673
6.1172 | -101.8467
6.121 | -101.8247
6.1247 | -101.8064
6.1285 | -101.7879
6.1322 | -101.7743
6.136 | -101.7716
6.1397 | -101.7733
6.1435 | -101.7848
6.1472 | -101.8051
6.151 | -101.8202
6.1547 | -101.8403
6.1585 | -101.8663
6.1622 | -101.8977
6.166 | -101.9284
6.1697 | -101.9531
6.1735 | -101.9782
6.1772 | -102.003
6.181 | -102.031
6.1847 | -102.0556
6.1885 | -102.071
6.1922 | -102.0826
6.196 | -102.0874
6.1997 | -102.0865
6.2035 | -102.0869
6.2072 | -102.0847
6.211 | -102.0748
6.2147 | -102.0648
6.2185 | -102.0575
6.2222 | -102.0543
6.226 | -102.0524
6.2297 | -102.0494
6.2335 | -102.0527
6.2372 | -102.0533
6.241 | -102.0556
6.2447 | -102.0643
6.2485 | -102.0729
6.2522 | -102.0797
6.256 | -102.0876
6.2597 | -102.0973
6.2635 | -102.107
6.2672 | -102.1155
6.271 | -102.124
6.2747 | -102.129
6.2785 | -102.1326
6.2822 | -102.1388
6.286 | -102.1445
6.2897 | -102.1564
6.2935 | -102.1714
6.2972 | -102.1899
6.301 | -102.221
6.3047 | -102.2618
6.3085 | -102.3022
6.3122 | -102.3367
6.316 | -102.3651
6.3197 | -102.3866
6.3235 | -102.3983
6.3272 | -102.3983
6.331 | -102.3837
6.3347 | -102.3569
6.3385 | -102.3215
6.3422 | -102.2834
6.346 | -102.2424
6.3497 | -102.2026
6.3535 | -102.1656
6.3572 | -102.1377
6.361 | -102.116
6.3647 | -102.1038
6.3685 | -102.1071
6.3722 | -102.1284
6.376 | -102.1627
6.3797 | -102.2087
6.3835 | -102.2625
6.3872 | -102.3261
6.391 | -102.393
6.3947 | -102.4536
6.3985 | -102.4946
6.4023 | -102.5188
6.406 | -102.5292
6.4098 | -102.527
6.4135 | -102.5074
6.4173 | -102.4759
6.421 | -102.4384
6.4248 | -102.4029
6.4285 | -102.3707
6.4323 | -102.3415
6.436 | -102.3174
6.4398 | -102.3021
6.4435 | -102.2968
6.4473 | -102.2983
6.451 | -102.3062
6.4548 | -102.3202
6.4585 | -102.3302
6.4623 | -102.3345
6.466 | -102.33
6.4698 | -102.3146
6.4735 | -102.2838
6.4773 | -102.2427
6.481 | -102.197
6.4848 | -102.1486
6.4885 | -102.1017
6.4923 | -102.0899
6.496 | -102.0717
6.4998 | -102.0659
6.5035 | -102.0689
6.5073 | -102.0811
6.511 | -102.1048
6.5148 | -102.1333
6.5185 | -102.1691
6.5223 | -102.2076
6.526 | -102.2432
6.5298 | -102.2842
6.5335 | -102.3274
6.5373 | -102.368
6.541 | -102.4042
6.5448 | -102.434
6.5485 | -102.4616
6.5523 | -102.4869
6.556 | -102.5099
6.5598 | -102.5291
6.5635 | -102.5502
6.5673 | -102.5667
6.571 | -102.5871
6.5748 | -102.6076
6.5785 | -102.6236
6.5823 | -102.64
6.586 | -102.6588
6.5898 | -102.6772
6.5935 | -102.6971
6.5973 | -102.7139
6.601 | -102.7218
6.6048 | -102.7284
6.6085 | -102.7371
6.6123 | -102.7435
6.616 | -102.7496
6.6198 | -102.7484
6.6235 | -102.7418
6.6273 | -102.7347
6.631 | -102.7234
6.6348 | -102.7121
6.6385 | -102.7069
6.6423 | -102.7049
6.646 | -102.71
6.6498 | -102.7205
6.6535 | -102.739
6.6573 | -102.7643
6.661 | -102.795
6.6648 | -102.8237
6.6685 | -102.8486
6.6723 | -102.8643
6.676 | -102.8695
6.6798 | -102.8685
6.6835 | -102.8552
6.6873 | -102.8261
6.691 | -102.7888
6.6948 | -102.7423
6.6985 | -102.6934
6.7023 | -102.6304
6.706 | -102.5631
6.7098 | -102.4931
6.7135 | -102.4201
6.7173 | -102.3455
6.721 | -102.2754
6.7248 | -102.2106
6.7286 | -102.1491
6.7323 | -102.0939
6.7361 | -102.0456
6.7398 | -102.0037
6.7436 | -101.971
6.7473 | -101.9463
6.7511 | -101.9293
6.7548 | -101.9247
6.7586 | -101.9317
6.7623 | -101.9455
6.7661 | -101.9689
6.7698 | -102.002
6.7736 | -102.0424
6.7773 | -102.0863
6.7811 | -102.1259
6.7848 | -102.1697
6.7886 | -102.2145
6.7923 | -102.2552
6.7961 | -102.2989
6.7998 | -102.3434
6.8036 | -102.3954
6.8073 | -102.4476
6.8111 | -102.5056
6.8148 | -102.5679
6.8186 | -102.6323
6.8223 | -102.7
6.8261 | -102.77
6.8298 | -102.8401
6.8336 | -102.9085
6.8373 | -102.9702
6.8411 | -103.0213
6.8448 | -103.0683
6.8486 | -103.1058
6.8523 | -103.1273
6.8561 | -103.1327
6.8598 | -103.1233
6.8636 | -103.1012
6.8673 | -103.0706
6.8711 | -103.0298
6.8748 | -102.9772
6.8786 | -102.9249
6.8823 | -102.8814
6.8861 | -102.8382
6.8898 | -102.7976
6.8936 | -102.766
6.8973 | -102.7392
6.9011 | -102.7135
6.9048 | -102.6921
6.9086 | -102.6677
6.9123 | -102.6453
6.9161 | -102.6289
6.9198 | -102.6114
6.9236 | -102.5922
6.9273 | -102.5722
6.9311 | -102.5504
6.9348 | -102.5325
6.9386 | -102.5151
6.9423 | -102.5011
6.9461 | -102.4856
6.9498 | -102.4677
6.9536 | -102.4547
6.9573 | -102.4438
6.9611 | -102.4333
6.9648 | -102.422
6.9686 | -102.4072
6.9723 | -102.3916
6.9761 | -102.3791
6.9798 | -102.3676
6.9836 | -102.3551
6.9873 | -102.3512
6.9911 | -102.3568
6.9948 | -102.3681
6.9986 | -102.3832
7.0023 | -102.4101
7.0061 | -102.4413
7.0098 | -102.4796
7.0136 | -102.5114
7.0173 | -102.5354
7.0211 | -102.5514
7.0248 | -102.5552
7.0286 | -102.5523
7.0323 | -102.5157
7.0361 | -102.4948
7.0399 | -102.4673
7.0436 | -102.4402
7.0474 | -102.4093
7.0511 | -102.3821
7.0549 | -102.3641
7.0586 | -102.3509
7.0624 | -102.341
7.0661 | -102.3457
7.0699 | -102.3586
7.0736 | -102.3841
7.0774 | -102.4206
7.0811 | -102.4672
7.0849 | -102.52
7.0886 | -102.576
7.0924 | -102.627
7.0961 | -102.6769
7.0999 | -102.7139
7.1036 | -102.7468
7.1074 | -102.7715
7.1111 | -102.788
7.1149 | -102.7937
7.1186 | -102.7947
7.1224 | -102.794
7.1261 | -102.79
7.1299 | -102.7842
7.1336 | -102.7759
7.1374 | -102.7657
7.1411 | -102.757
7.1449 | -102.7562
7.1486 | -102.7559
7.1524 | -102.7555
7.1561 | -102.7552
7.1599 | -102.7539
7.1636 | -102.7552
7.1674 | -102.7521
7.1711 | -102.7546
7.1749 | -102.7531
7.1786 | -102.7528
7.1824 | -102.7434
7.1861 | -102.7366
7.1899 | -102.7255
7.1936 | -102.7136
7.1974 | -102.7096
7.2011 | -102.706
7.2049 | -102.6978
7.2086 | -102.6971
7.2124 | -102.695
7.2161 | -102.6892
7.2199 | -102.687
7.2236 | -102.6879
7.2274 | -102.695
7.2311 | -102.7028
7.2349 | -102.7192
7.2386 | -102.7387
7.2424 | -102.7634
7.2461 | -102.7975
7.2499 | -102.8342
7.2536 | -102.8695
7.2574 | -102.9049
7.2611 | -102.9509
7.2649 | -102.9991
7.2686 | -103.0469
7.2724 | -103.0958
7.2761 | -103.1482
7.2799 | -103.1889
7.2836 | -103.2264
7.2874 | -103.2601
7.2911 | -103.2893
7.2949 | -103.3073
7.2986 | -103.3195
7.3024 | -103.3222
7.3061 | -103.3163
7.3099 | -103.3095
7.3136 | -103.2962
7.3174 | -103.2772
7.3211 | -103.2613
7.3249 | -103.2389
7.3286 | -103.2115
7.3324 | -103.1899
7.3361 | -103.1638
7.3399 | -103.1405
7.3436 | -103.1153
7.3474 | -103.0845
7.3511 | -103.0481
7.3549 | -103.0146
7.3586 | -102.9857
7.3624 | -102.9592
7.3662 | -102.933
7.3699 | -102.9055
7.3737 | -102.8828
7.3774 | -102.8659
7.3812 | -102.8552
7.3849 | -102.8509
7.3887 | -102.854
7.3924 | -102.8627
7.3962 | -102.8789
7.3999 | -102.9034
7.4037 | -102.928
7.4074 | -102.9592
7.4112 | -102.9952
7.4149 | -103.0274
7.4187 | -103.0499
7.4224 | -103.0688
7.4262 | -103.083
7.4299 | -103.0823
7.4337 | -103.0792
7.4374 | -103.0752
7.4412 | -103.0648
7.4449 | -103.0565
7.4487 | -103.0487
7.4524 | -103.0411
7.4562 | -103.0411
7.4599 | -103.0443
7.4637 | -103.0394
7.4674 | -103.0364
7.4712 | -103.0438
7.4749 | -103.0563
7.4787 | -103.0701
7.4824 | -103.092
7.4862 | -103.113
7.4899 | -103.1356
7.4937 | -103.1623
7.4974 | -103.1978
7.5012 | -103.2301
7.5049 | -103.2579
7.5087 | -103.2809
7.5124 | -103.3025
7.5162 | -103.3154
7.5199 | -103.3231
7.5237 | -103.3279
7.5274 | -103.3376
7.5312 | -103.3478
7.5349 | -103.3489
7.5387 | -103.355
7.5424 | -103.3658
7.5462 | -103.3699
7.5499 | -103.3731
7.5537 | -103.3801
7.5574 | -103.3853
7.5612 | -103.3893
7.5649 | -103.3794
7.5687 | -103.3564
7.5724 | -103.3058
7.5762 | -103.2797
7.5799 | -103.2436
7.5837 | -103.2013
7.5874 | -103.1566
7.5912 | -103.116
7.5949 | -103.0799
7.5987 | -103.0518
7.6024 | -103.0314
7.6062 | -103.0154
7.6099 | -103.0019
7.6137 | -102.9988
7.6174 | -103.004
7.6212 | -103.0183
7.6249 | -103.0375
7.6287 | -103.0576
7.6324 | -103.0794
7.6362 | -103.1051
7.6399 | -103.1319
7.6437 | -103.1544
7.6474 | -103.184
7.6512 | -103.2107
7.6549 | -103.2321
7.6587 | -103.2508
7.6624 | -103.2644
7.6662 | -103.2671
7.6699 | -103.2603
7.6737 | -103.2485
7.6775 | -103.243
7.6812 | -103.2342
7.685 | -103.2195
7.6887 | -103.2028
7.6925 | -103.1773
7.6962 | -103.1525
7.7 | -103.1341
7.7037 | -103.1195
7.7075 | -103.1014
7.7112 | -103.0864
7.715 | -103.0739
7.7187 | -103.0573
7.7225 | -103.041
7.7262 | -103.0317
7.73 | -103.0224
7.7337 | -103.013
7.7375 | -103.0095
7.7412 | -103.0105
7.745 | -103.0128
7.7487 | -103.0176
7.7525 | -103.0247
7.7562 | -103.0328
7.76 | -103.0407
7.7637 | -103.0494
7.7675 | -103.059
7.7712 | -103.0714
7.775 | -103.085
7.7787 | -103.0961
7.7825 | -103.1061
7.7862 | -103.1139
7.79 | -103.1159
7.7937 | -103.1134
7.7975 | -103.1136
7.8012 | -103.1176
7.805 | -103.1254
7.8087 | -103.1394
7.8125 | -103.1573
7.8162 | -103.1702
7.82 | -103.1893
7.8237 | -103.2175
7.8275 | -103.2416
7.8312 | -103.265
7.835 | -103.2947
7.8387 | -103.3291
7.8425 | -103.356
7.8462 | -103.3829
7.85 | -103.406
7.8537 | -103.421
7.8575 | -103.4326
7.8612 | -103.4466
7.865 | -103.4524
7.8687 | -103.4472
7.8725 | -103.4381
7.8762 | -103.4257
7.88 | -103.4075
7.8837 | -103.3855
7.8875 | -103.3614
7.8912 | -103.3408
7.895 | -103.3223
7.8987 | -103.3097
7.9025 | -103.3002
7.9062 | -103.2954
7.91 | -103.2919
7.9137 | -103.292
7.9175 | -103.2902
7.9212 | -103.2959
7.925 | -103.3052
7.9287 | -103.3158
7.9325 | -103.3337
7.9362 | -103.3554
7.94 | -103.3759
7.9437 | -103.4041
7.9475 | -103.4382
7.9512 | -103.4679
7.955 | -103.4902
7.9587 | -103.5169
7.9625 | -103.543
7.9662 | -103.5642
7.97 | -103.5722
7.9737 | -103.5756
7.9775 | -103.5694
7.9812 | -103.5595
7.985 | -103.5441
7.9887 | -103.5189
7.9925 | -103.4863
7.9962 | -103.4482
8 | -103.4051
8.0038 | -103.3527
8.0075 | -103.3031
8.0113 | -103.2518
8.015 | -103.2001
8.0188 | -103.157
8.0225 | -103.1162
8.0263 | -103.0824
8.03 | -103.0554
8.0338 | -103.0417
8.0375 | -103.0353
8.0413 | -103.045
8.045 | -103.0707
8.0488 | -103.1061
8.0525 | -103.1474
8.0563 | -103.1957
8.06 | -103.2497
8.0638 | -103.3038
8.0675 | -103.3599
8.0713 | -103.4149
8.075 | -103.4591
8.0788 | -103.496
8.0825 | -103.526
8.0863 | -103.5517
8.09 | -103.5619
8.0938 | -103.5564
8.0975 | -103.5397
8.1013 | -103.5181
8.105 | -103.4907
8.1088 | -103.4628
8.1125 | -103.4246
8.1163 | -103.3769
8.12 | -103.3389
8.1238 | -103.3014
8.1275 | -103.2643
8.1313 | -103.2319
8.135 | -103.1998
8.1388 | -103.1727
8.1425 | -103.1557
8.1463 | -103.1488
8.15 | -103.1466
8.1538 | -103.1497
8.1575 | -103.1606
8.1613 | -103.1731
8.165 | -103.1899
8.1688 | -103.2101
8.1725 | -103.237
8.1763 | -103.2677
8.18 | -103.2968
8.1838 | -103.3292
8.1875 | -103.3578
8.1913 | -103.3876
8.195 | -103.415
8.1988 | -103.4394
8.2025 | -103.466
8.2063 | -103.4881
8.21 | -103.5085
8.2138 | -103.5327
8.2175 | -103.5526
8.2213 | -103.566
8.225 | -103.5797
8.2288 | -103.5874
8.2325 | -103.5985
8.2363 | -103.61
8.24 | -103.6153
8.2438 | -103.6163
8.2475 | -103.6164
8.2513 | -103.6156
8.255 | -103.6178
8.2588 | -103.6238
8.2625 | -103.6263
8.2663 | -103.6291
8.27 | -103.6347
8.2738 | -103.6465
8.2775 | -103.6593
8.2813 | -103.6714
8.285 | -103.6842
8.2888 | -103.6984
8.2925 | -103.7169
8.2963 | -103.7373
8.3 | -103.7608
8.3038 | -103.7846
8.3075 | -103.8095
8.3113 | -103.8393
8.315 | -103.8653
8.3188 | -103.8918
8.3225 | -103.9223
8.3263 | -103.951
8.3301 | -103.9733
8.3338 | -103.9922
8.3376 | -104.0106
8.3413 | -104.031
8.3451 | -104.0483
8.3488 | -104.0546
8.3526 | -104.0556
8.3563 | -104.0552
8.3601 | -104.0548
8.3638 | -104.0489
8.3676 | -104.0379
8.3713 | -104.0223
8.3751 | -103.998
8.3788 | -103.9688
8.3826 | -103.941
8.3863 | -103.9131
8.3901 | -103.8764
8.3938 | -103.8399
8.3976 | -103.8041
8.4013 | -103.7708
8.4051 | -103.7378
8.4088 | -103.7023
8.4126 | -103.6681
8.4163 | -103.6314
8.4201 | -103.5942
8.4238 | -103.5608
8.4276 | -103.5302
8.4313 | -103.4989
8.4351 | -103.4693
8.4388 | -103.447
8.4426 | -103.4313
8.4463 | -103.4159
8.4501 | -103.4056
8.4538 | -103.4056
8.4576 | -103.4134
8.4613 | -103.4311
8.4651 | -103.4588
8.4688 | -103.4973
8.4726 | -103.5436
8.4763 | -103.5982
8.4801 | -103.6523
8.4838 | -103.7069
8.4876 | -103.7672
8.4913 | -103.8254
8.4951 | -103.8825
8.4988 | -103.933
8.5026 | -103.9789
8.5063 | -104.0199
8.5101 | -104.0545
8.5138 | -104.0848
8.5176 | -104.1036
8.5213 | -104.1196
8.5251 | -104.1317
8.5288 | -104.1415
8.5326 | -104.1484
8.5363 | -104.1467
8.5401 | -104.146
8.5438 | -104.1397
8.5476 | -104.1303
8.5513 | -104.1241
8.5551 | -104.1209
8.5588 | -104.1134
8.5626 | -104.1098
8.5663 | -104.1076
8.5701 | -104.1082
8.5738 | -104.1101
8.5776 | -104.1153
8.5813 | -104.1193
8.5851 | -104.1226
8.5888 | -104.1283
8.5926 | -104.1388
8.5963 | -104.1521
8.6001 | -104.1611
8.6038 | -104.1673
8.6076 | -104.1714
8.6113 | -104.1796
8.6151 | -104.1927
8.6188 | -104.202
8.6226 | -104.2138
8.6263 | -104.2274
8.6301 | -104.2347
8.6338 | -104.2394
8.6376 | -104.2461
8.6413 | -104.2455
8.6451 | -104.2421
8.6489 | -104.2337
8.6526 | -104.2214
8.6564 | -104.1908
8.6601 | -104.1706
8.6639 | -104.1464
8.6676 | -104.1169
8.6714 | -104.0913
8.6751 | -104.0604
8.6789 | -104.0295
8.6826 | -104.002
8.6864 | -103.9777
8.6901 | -103.9509
8.6939 | -103.9324
8.6976 | -103.9205
8.7014 | -103.9166
8.7051 | -103.9162
8.7089 | -103.9207
8.7126 | -103.9299
8.7164 | -103.9412
8.7201 | -103.9533
8.7239 | -103.9716
8.7276 | -103.9941
8.7314 | -104.0153
8.7351 | -104.0326
8.7389 | -104.0438
8.7426 | -104.0508
8.7464 | -104.0607
8.7501 | -104.0718
8.7539 | -104.0799
8.7576 | -104.0892
8.7614 | -104.095
8.7651 | -104.1012
8.7689 | -104.1114
8.7726 | -104.1157
8.7764 | -104.1225
8.7801 | -104.1332
8.7839 | -104.1461
8.7876 | -104.1571
8.7914 | -104.174
8.7951 | -104.1941
8.7989 | -104.215
8.8026 | -104.2376
8.8064 | -104.2627
8.8101 | -104.2834
8.8139 | -104.2998
8.8176 | -104.3132
8.8214 | -104.3234
8.8251 | -104.3261
8.8289 | -104.3225
8.8326 | -104.3117
8.8364 | -104.2978
8.8401 | -104.2754
8.8439 | -104.243
8.8476 | -104.2061
8.8514 | -104.1693
8.8551 | -104.1325
8.8589 | -104.0937
8.8626 | -104.0569
8.8664 | -104.0182
8.8701 | -103.9799
8.8739 | -103.9436
8.8776 | -103.9074
8.8814 | -103.8699
8.8851 | -103.8356
8.8889 | -103.8104
8.8926 | -103.7945
8.8964 | -103.78
8.9001 | -103.7687
8.9039 | -103.7635
8.9076 | -103.7651
8.9114 | -103.7704
8.9151 | -103.7802
8.9189 | -103.7983
8.9226 | -103.8183
8.9264 | -103.8408
8.9301 | -103.8653
8.9339 | -103.8886
8.9376 | -103.9152
8.9414 | -103.944
8.9451 | -103.9737
8.9489 | -104.0019
8.9526 | -104.0261
8.9564 | -104.0475
8.9602 | -104.0639
8.9639 | -104.0816
8.9677 | -104.0988
8.9714 | -104.1128
8.9752 | -104.1261
8.9789 | -104.1382
8.9827 | -104.1439
8.9864 | -104.1392
8.9902 | -104.1345
8.9939 | -104.1272
8.9977 | -104.1181
9.0014 | -104.1023
9.0052 | -104.0749
9.0089 | -104.0473
9.0127 | -104.0258
9.0164 | -104.0057
9.0202 | -103.9803
9.0239 | -103.9544
9.0277 | -103.9356
9.0314 | -103.9207
9.0352 | -103.9107
9.0389 | -103.9044
9.0427 | -103.9052
9.0464 | -103.9104
9.0502 | -103.9191
9.0539 | -103.9329
9.0577 | -103.9553
9.0614 | -103.9825
9.0652 | -104.0126
9.0689 | -104.0425
9.0727 | -104.0741
9.0764 | -104.109
9.0802 | -104.1431
9.0839 | -104.176
9.0877 | -104.2079
9.0914 | -104.2398
9.0952 | -104.2646
9.0989 | -104.2862
9.1027 | -104.3103
9.1064 | -104.3343
9.1102 | -104.3566
9.1139 | -104.3799
9.1177 | -104.4067
9.1214 | -104.4273
9.1252 | -104.4544
9.1289 | -104.4852
9.1327 | -104.5167
9.1364 | -104.5551
9.1402 | -104.5849
9.1439 | -104.6014
9.1477 | -104.6138
9.1514 | -104.626
9.1552 | -104.6386
9.1589 | -104.645
9.1627 | -104.6438
9.1664 | -104.636
9.1702 | -104.6259
9.1739 | -104.6112
9.1777 | -104.5959
9.1814 | -104.577
9.1852 | -104.5595
9.1889 | -104.5458
9.1927 | -104.5344
9.1964 | -104.5433
9.2002 | -104.5415
9.2039 | -104.5448
9.2077 | -104.5474
9.2114 | -104.5495
9.2152 | -104.5553
9.2189 | -104.5668
9.2227 | -104.5758
9.2264 | -104.5799
9.2302 | -104.5821
9.2339 | -104.5828
9.2377 | -104.5746
9.2414 | -104.5608
9.2452 | -104.5462
9.2489 | -104.5266
9.2527 | -104.5027
9.2564 | -104.4776
9.2602 | -104.4413
9.2639 | -104.4033
9.2677 | -104.3687
9.2714 | -104.3341
9.2752 | -104.2995
9.279 | -104.2644
9.2827 | -104.2328
9.2865 | -104.2006
9.2902 | -104.1669
9.294 | -104.1355
9.2977 | -104.1179
9.3015 | -104.1069
9.3052 | -104.0971
9.309 | -104.0926
9.3127 | -104.0928
9.3165 | -104.0984
9.3202 | -104.1087
9.324 | -104.1212
9.3277 | -104.142
9.3315 | -104.1669
9.3352 | -104.1919
9.339 | -104.2184
9.3427 | -104.2465
9.3465 | -104.2771
9.3502 | -104.3056
9.354 | -104.3331
9.3577 | -104.3554
9.3615 | -104.374
9.3652 | -104.3991
9.369 | -104.4252
9.3727 | -104.4478
9.3765 | -104.4715
9.3802 | -104.4908
9.384 | -104.5065
9.3877 | -104.5241
9.3915 | -104.5374
9.3952 | -104.5475
9.399 | -104.5602
9.4027 | -104.571
9.4065 | -104.5751
9.4102 | -104.579
9.414 | -104.5794
9.4177 | -104.5816
9.4215 | -104.5764
9.4252 | -104.5725
9.429 | -104.5708
9.4327 | -104.5618
9.4365 | -104.5454
9.4402 | -104.526
9.444 | -104.5051
9.4477 | -104.4824
9.4515 | -104.4592
9.4552 | -104.4313
9.459 | -104.401
9.4627 | -104.3738
9.4665 | -104.3492
9.4702 | -104.3272
9.474 | -104.3088
9.4777 | -104.2961
9.4815 | -104.2859
9.4852 | -104.2811
9.489 | -104.2806
9.4927 | -104.2809
9.4965 | -104.2849
9.5002 | -104.2941
9.504 | -104.302
9.5077 | -104.3106
9.5115 | -104.3217
9.5152 | -104.3334
9.519 | -104.3436
9.5227 | -104.3539
9.5265 | -104.3623
9.5302 | -104.3746
9.534 | -104.3817
9.5377 | -104.3864
9.5415 | -104.3902
9.5452 | -104.3932
9.549 | -104.3929
9.5527 | -104.3946
9.5565 | -104.3946
9.5602 | -104.392
9.564 | -104.3873
9.5677 | -104.3778
9.5715 | -104.3728
9.5752 | -104.3633
9.579 | -104.3512
9.5827 | -104.3335
9.5865 | -104.3084
9.5902 | -104.2864
9.594 | -104.262
9.5977 | -104.2347
9.6015 | -104.204
9.6053 | -104.1738
9.609 | -104.1457
9.6128 | -104.1214
9.6165 | -104.1013
9.6203 | -104.0841
9.624 | -104.0716
9.6278 | -104.0657
9.6315 | -104.0659
9.6353 | -104.072
9.639 | -104.0855
9.6428 | -104.1054
9.6465 | -104.1288
9.6503 | -104.155
9.654 | -104.1889
9.6578 | -104.2262
9.6615 | -104.261
9.6653 | -104.2954
9.669 | -104.333
9.6728 | -104.3746
9.6765 | -104.4101
9.6803 | -104.4397
9.684 | -104.4695
9.6878 | -104.5008
9.6915 | -104.5279
9.6953 | -104.5525
9.699 | -104.5744
9.7028 | -104.5926
9.7065 | -104.612
9.7103 | -104.6248
9.714 | -104.6393
9.7178 | -104.6574
9.7215 | -104.6765
9.7253 | -104.6975
9.729 | -104.7145
9.7328 | -104.7362
9.7365 | -104.7575
9.7403 | -104.7786
9.744 | -104.7978
9.7478 | -104.8167
9.7515 | -104.8359
9.7553 | -104.8504
9.759 | -104.863
9.7628 | -104.8737
9.7665 | -104.8789
9.7703 | -104.8783
9.774 | -104.8786
9.7778 | -104.8734
9.7815 | -104.8643
9.7853 | -104.8501
9.789 | -104.8335
9.7928 | -104.8148
9.7965 | -104.7882
9.8003 | -104.7631
9.804 | -104.7398
9.8078 | -104.7121
9.8115 | -104.6777
9.8153 | -104.6425
9.819 | -104.6106
9.8228 | -104.5832
9.8265 | -104.5573
9.8303 | -104.5295
9.834 | -104.5029
9.8378 | -104.4804
9.8415 | -104.4618
9.8453 | -104.4473
9.849 | -104.4366
9.8528 | -104.4335
9.8565 | -104.432
9.8603 | -104.4321
9.864 | -104.4405
9.8678 | -104.4493
9.8715 | -104.4608
9.8753 | -104.4782
9.879 | -104.4994
9.8828 | -104.5199
9.8865 | -104.5391
9.8903 | -104.5656
9.894 | -104.5906
9.8978 | -104.6158
9.9015 | -104.644
9.9053 | -104.6751
9.909 | -104.7069
9.9128 | -104.7368
9.9165 | -104.7683
9.9203 | -104.7976
9.9241 | -104.8297
9.9278 | -104.8613
9.9316 | -104.8929
9.9353 | -104.9263
9.9391 | -104.9527
9.9428 | -104.977
9.9466 | -104.9996
9.9503 | -105.0167
9.9541 | -105.0308
9.9578 | -105.0419
9.9616 | -105.0496
9.9653 | -105.0538
9.9691 | -105.0516
9.9728 | -105.0426
9.9766 | -105.0346
9.9803 | -105.0266
9.9841 | -105.018
9.9878 | -105.0089
9.9916 | -104.9947
9.9953 | -104.9793
9.9991 | -104.9651
10.0028 | -104.9536
10.0066 | -104.9413
10.0103 | -104.9251
10.0141 | -104.9129
10.0178 | -104.8998
10.0216 | -104.8857
10.0253 | -104.8766
10.0291 | -104.8686
10.0328 | -104.8651
10.0366 | -104.8599
10.0403 | -104.8541
10.0441 | -104.8524
10.0478 | -104.8502
10.0516 | -104.843
10.0553 | -104.8362
10.0591 | -104.831
10.0628 | -104.8216
10.0666 | -104.8094
10.0703 | -104.7945
10.0741 | -104.7815
10.0778 | -104.7668
10.0816 | -104.7487
10.0853 | -104.7377
10.0891 | -104.7259
10.0928 | -104.7093
10.0966 | -104.696
10.1003 | -104.6849
10.1041 | -104.6753
10.1078 | -104.6703
10.1116 | -104.669
10.1153 | -104.672
10.1191 | -104.681
10.1228 | -104.6896
10.1266 | -104.7039
10.1303 | -104.7208
10.1341 | -104.7414
10.1378 | -104.7674
10.1416 | -104.796
10.1453 | -104.8222
10.1491 | -104.8464
10.1528 | -104.873
10.1566 | -104.9011
10.1603 | -104.9238
10.1641 | -104.9422
10.1678 | -104.9545
10.1716 | -104.9587
10.1753 | -104.9552
10.1791 | -104.9444
10.1828 | -104.9329
10.1866 | -104.9182
10.1903 | -104.894
10.1941 | -104.8672
10.1978 | -104.8391
10.2016 | -104.8094
10.2053 | -104.7778
10.2091 | -104.7504
10.2128 | -104.7192
10.2166 | -104.687
10.2203 | -104.6544
10.2241 | -104.6277
10.2278 | -104.6084
10.2316 | -104.5921
10.2353 | -104.5806
10.2391 | -104.5731
10.2429 | -104.5725
10.2466 | -104.5775
10.2504 | -104.5855
10.2541 | -104.5971
10.2579 | -104.6113
10.2616 | -104.6316
10.2654 | -104.656
10.2691 | -104.6812
10.2729 | -104.7093
10.2766 | -104.7395
10.2804 | -104.771
10.2841 | -104.7978
10.2879 | -104.8263
10.2916 | -104.8532
10.2954 | -104.8753
10.2991 | -104.8959
10.3029 | -104.9094
10.3066 | -104.9196
10.3104 | -104.9304
10.3141 | -104.9316
10.3179 | -104.9319
10.3216 | -104.9308
10.3254 | -104.9266
10.3291 | -104.9153
10.3329 | -104.8946
10.3366 | -104.8759
10.3404 | -104.8587
10.3441 | -104.8434
10.3479 | -104.8264
10.3516 | -104.8028
10.3554 | -104.7824
10.3591 | -104.7607
10.3629 | -104.7451
10.3666 | -104.7344
10.3704 | -104.7227
10.3741 | -104.7059
10.3779 | -104.6899
10.3816 | -104.6762
10.3854 | -104.6634
10.3891 | -104.6555
10.3929 | -104.6481
10.3966 | -104.6442
10.4004 | -104.6406
10.4041 | -104.6346
10.4079 | -104.6309
10.4116 | -104.6241
10.4154 | -104.6194
10.4191 | -104.6222
10.4229 | -104.6191
10.4266 | -104.6155
10.4304 | -104.6134
10.4341 | -104.6161
10.4379 | -104.6174
10.4416 | -104.616
10.4454 | -104.6144
10.4491 | -104.6158
10.4529 | -104.6185
10.4566 | -104.6258
10.4604 | -104.6305
10.4641 | -104.6346
10.4679 | -104.6448
10.4716 | -104.6541
10.4754 | -104.6655
10.4791 | -104.6833
10.4829 | -104.6994
10.4866 | -104.7198
10.4904 | -104.7396
10.4941 | -104.7602
10.4979 | -104.7764
10.5016 | -104.7943
10.5054 | -104.816
10.5091 | -104.8304
10.5129 | -104.8427
10.5166 | -104.852
10.5204 | -104.8563
10.5241 | -104.8593
10.5279 | -104.8621
10.5316 | -104.8575
10.5354 | -104.8539
10.5391 | -104.8483
10.5429 | -104.8355
10.5466 | -104.8246
10.5504 | -104.8137
10.5541 | -104.8056
10.5579 | -104.7992
10.5617 | -104.7932
10.5654 | -104.7908
10.5692 | -104.7856
10.5729 | -104.7758
10.5767 | -104.7729
10.5804 | -104.7713
10.5842 | -104.7725
10.5879 | -104.776
10.5917 | -104.7814
10.5954 | -104.7869
10.5992 | -104.7942
10.6029 | -104.7995
10.6067 | -104.802
10.6104 | -104.8037
10.6142 | -104.8054
10.6179 | -104.8065
10.6217 | -104.8042
10.6254 | -104.797
10.6292 | -104.786
10.6329 | -104.7679
10.6367 | -104.7503
10.6404 | -104.7272
10.6442 | -104.6999
10.6479 | -104.675
10.6517 | -104.6498
10.6554 | -104.6214
10.6592 | -104.594
10.6629 | -104.5691
10.6667 | -104.5398
10.6704 | -104.512
10.6742 | -104.4884
10.6779 | -104.4683
10.6817 | -104.4518
10.6854 | -104.4428
10.6892 | -104.4347
10.6929 | -104.4347
10.6967 | -104.4405
10.7004 | -104.4484
10.7042 | -104.465
10.7079 | -104.4845
10.7117 | -104.5075
10.7154 | -104.5348
10.7192 | -104.568
10.7229 | -104.6065
10.7267 | -104.6454
10.7304 | -104.6819
10.7342 | -104.716
10.7379 | -104.7472
10.7417 | -104.7736
10.7454 | -104.7967
10.7492 | -104.815
10.7529 | -104.8231
10.7567 | -104.8273
10.7604 | -104.8275
10.7642 | -104.821
10.7679 | -104.812
10.7717 | -104.7993
10.7754 | -104.7839
10.7792 | -104.7649
10.7829 | -104.7435
10.7867 | -104.7201
10.7904 | -104.6964
10.7942 | -104.6734
10.7979 | -104.6515
10.8017 | -104.6257
10.8054 | -104.6059
10.8092 | -104.5921
10.8129 | -104.5808
10.8167 | -104.5724
10.8204 | -104.5834
10.8242 | -104.586
10.8279 | -104.5901
10.8317 | -104.5946
10.8354 | -104.6012
10.8392 | -104.6099
10.8429 | -104.6208
10.8467 | -104.6335
10.8504 | -104.6457
10.8542 | -104.6633
10.8579 | -104.6849
10.8617 | -104.7032
10.8654 | -104.7246
10.8692 | -104.747
10.8729 | -104.7683
10.8767 | -104.7832
10.8804 | -104.7955
10.8842 | -104.8138
10.888 | -104.8327
10.8917 | -104.8543
10.8955 | -104.8721
10.8992 | -104.8886
10.903 | -104.9023
10.9067 | -104.9177
10.9105 | -104.9332
10.9142 | -104.9443
10.918 | -104.9535
10.9217 | -104.9625
10.9255 | -104.9662
10.9292 | -104.9707
10.933 | -104.9794
10.9367 | -104.9866
10.9405 | -104.9941
10.9442 | -105.0023
10.948 | -105.0048
10.9517 | -105.0037
10.9555 | -105.0029
10.9592 | -105.0027
10.963 | -104.9992
10.9667 | -104.9913
10.9705 | -104.9801
10.9742 | -104.9683
10.978 | -104.9519
10.9817 | -104.9367
10.9855 | -104.9185
10.9892 | -104.8992
10.993 | -104.8793
10.9967 | -104.8587
11.0005 | -104.8368
11.0042 | -104.8142
11.008 | -104.794
11.0117 | -104.7787
11.0155 | -104.7634
11.0192 | -104.7492
11.023 | -104.7425
11.0267 | -104.7399
11.0305 | -104.7419
11.0342 | -104.7456
11.038 | -104.7569
11.0417 | -104.7707
11.0455 | -104.7808
11.0492 | -104.7957
11.053 | -104.8123
11.0567 | -104.8304
11.0605 | -104.8502
11.0642 | -104.872
11.068 | -104.8878
11.0717 | -104.9022
11.0755 | -104.9161
11.0792 | -104.923
11.083 | -104.9279
11.0867 | -104.9331
11.0905 | -104.928
11.0942 | -104.9168
11.098 | -104.9079
11.1017 | -104.8938
11.1055 | -104.8761
11.1092 | -104.8555
11.113 | -104.8321
11.1167 | -104.804
11.1205 | -104.7751
11.1242 | -104.7465
11.128 | -104.7211
11.1317 | -104.698
11.1355 | -104.6786
11.1392 | -104.66
11.143 | -104.6403
11.1467 | -104.6266
11.1505 | -104.6184
11.1542 | -104.6148
11.158 | -104.6141
11.1617 | -104.6193
11.1655 | -104.6255
11.1692 | -104.6327
11.173 | -104.6404
11.1767 | -104.6527
11.1805 | -104.6672
11.1842 | -104.6784
11.188 | -104.689
11.1917 | -104.7031
11.1955 | -104.7213
11.1992 | -104.7393
11.203 | -104.7555
11.2068 | -104.7728
11.2105 | -104.7854
11.2143 | -104.7964
11.218 | -104.8106
11.2218 | -104.8245
11.2255 | -104.8376
11.2293 | -104.853
11.233 | -104.8652
11.2368 | -104.8759
11.2405 | -104.8873
11.2443 | -104.8957
11.248 | -104.9038
11.2518 | -104.9128
11.2555 | -104.9191
11.2593 | -104.9247
11.263 | -104.9261
11.2668 | -104.9225
11.2705 | -104.9192
11.2743 | -104.9109
11.278 | -104.9004
11.2818 | -104.8847
11.2855 | -104.8646
11.2893 | -104.8411
11.293 | -104.8188
11.2968 | -104.7922
11.3005 | -104.7676
11.3043 | -104.7474
11.308 | -104.7246
11.3118 | -104.7016
11.3155 | -104.6838
11.3193 | -104.6674
11.323 | -104.6491
11.3268 | -104.64
11.3305 | -104.6386
11.3343 | -104.6382
11.338 | -104.641
11.3418 | -104.6472
11.3455 | -104.6588
11.3493 | -104.674
11.353 | -104.6859
11.3568 | -104.6989
11.3605 | -104.7169
11.3643 | -104.7306
11.368 | -104.7484
11.3718 | -104.765
11.3755 | -104.7823
11.3793 | -104.7967
11.383 | -104.8085
11.3868 | -104.819
11.3905 | -104.8267
11.3943 | -104.8362
11.398 | -104.8506
11.4018 | -104.8633
11.4055 | -104.8777
11.4093 | -104.8907
11.413 | -104.9046
11.4168 | -104.9186
11.4205 | -104.9313
11.4243 | -104.9424
11.428 | -104.9578
11.4318 | -104.9752
11.4355 | -104.9902
11.4393 | -105.0086
11.443 | -105.0258
11.4468 | -105.0389
11.4505 | -105.0527
11.4543 | -105.0675
11.458 | -105.0809
11.4618 | -105.0952
11.4655 | -105.1103
11.4693 | -105.1248
11.473 | -105.1402
11.4768 | -105.1561
11.4805 | -105.1716
11.4843 | -105.1902
11.488 | -105.2115
11.4918 | -105.2298
11.4955 | -105.2448
11.4993 | -105.2626
11.503 | -105.278
11.5068 | -105.2953
11.5105 | -105.3135
11.5143 | -105.3292
11.5181 | -105.3443
11.5218 | -105.3611
11.5256 | -105.3739
11.5293 | -105.3845
11.5331 | -105.3964
11.5368 | -105.4061
11.5406 | -105.4118
11.5443 | -105.4197
11.5481 | -105.4295
11.5518 | -105.4368
11.5556 | -105.446
11.5593 | -105.4548
11.5631 | -105.4549
11.5668 | -105.4503
11.5706 | -105.445
11.5743 | -105.4369
11.5781 | -105.4293
11.5818 | -105.4185
11.5856 | -105.4074
11.5893 | -105.3981
11.5931 | -105.3858
11.5968 | -105.3729
11.6006 | -105.3565
11.6043 | -105.3415
11.6081 | -105.3257
11.6118 | -105.3105
11.6156 | -105.2993
11.6193 | -105.2883
11.6231 | -105.2769
11.6268 | -105.2631
11.6306 | -105.2527
11.6343 | -105.2439
11.6381 | -105.2346
11.6418 | -105.2255
11.6456 | -105.2165
11.6493 | -105.2047
11.6531 | -105.1978
11.6568 | -105.1926
11.6606 | -105.188
11.6643 | -105.1847
11.6681 | -105.1823
11.6718 | -105.1801
11.6756 | -105.1726
11.6793 | -105.1656
11.6831 | -105.1603
11.6868 | -105.1547
11.6906 | -105.1457
11.6943 | -105.1397
11.6981 | -105.1335
11.7018 | -105.1248
11.7056 | -105.116
11.7093 | -105.1098
11.7131 | -105.0975
11.7168 | -105.082
11.7206 | -105.0694
11.7243 | -105.0542
11.7281 | -105.0432
11.7318 | -105.0335
11.7356 | -105.0255
11.7393 | -105.0225
11.7431 | -105.022
11.7468 | -105.0215
11.7506 | -105.0262
11.7543 | -105.0343
11.7581 | -105.0413
11.7618 | -105.0504
11.7656 | -105.0616
11.7693 | -105.0782
11.7731 | -105.092
11.7768 | -105.1083
11.7806 | -105.1253
11.7843 | -105.1414
11.7881 | -105.1569
11.7918 | -105.1739
11.7956 | -105.192
11.7993 | -105.2071
11.8031 | -105.2204
11.8068 | -105.2342
11.8106 | -105.2441
11.8143 | -105.2569
11.8181 | -105.2688
11.8218 | -105.2759
11.8256 | -105.2782
11.8293 | -105.2777
11.8331 | -105.2783
11.8369 | -105.2786
11.8406 | -105.2769
11.8444 | -105.2773
11.8481 | -105.275
11.8519 | -105.269
11.8556 | -105.2657
11.8594 | -105.2596
11.8631 | -105.2512
11.8669 | -105.2449
11.8706 | -105.2383
11.8744 | -105.23
11.8781 | -105.2232
11.8819 | -105.2164
11.8856 | -105.2139
11.8894 | -105.21
11.8931 | -105.2066
11.8969 | -105.2002
11.9006 | -105.2014
11.9044 | -105.2016
11.9081 | -105.2016
11.9119 | -105.1993
11.9156 | -105.2002
11.9194 | -105.2034
11.9231 | -105.2032
11.9269 | -105.1991
11.9306 | -105.2001
11.9344 | -105.1997
11.9381 | -105.191
11.9419 | -105.1873
11.9456 | -105.1868
11.9494 | -105.1794
11.9531 | -105.173
11.9569 | -105.1661
11.9606 | -105.1555
11.9644 | -105.1418
11.9681 | -105.1269
11.9719 | -105.1135
11.9756 | -105.0977
11.9794 | -105.078
11.9831 | -105.0618
11.9869 | -105.0422
11.9906 | -105.0226
11.9944 | -105.0037
11.9981 | -104.9914
12.0019 | -104.98
12.0056 | -104.9686
12.0094 | -104.961
12.0131 | -104.9584
12.0169 | -104.9589
12.0206 | -104.9616
12.0244 | -104.9689
12.0281 | -104.9778
12.0319 | -104.99
12.0356 | -105.0025
12.0394 | -105.0172
12.0431 | -105.0342
12.0469 | -105.0526
12.0506 | -105.0716
12.0544 | -105.0915
12.0581 | -105.1092
12.0619 | -105.124
12.0656 | -105.138
12.0694 | -105.1515
12.0731 | -105.1604
12.0769 | -105.1688
12.0806 | -105.174
12.0844 | -105.1784
12.0881 | -105.1806
12.0919 | -105.1786
12.0956 | -105.1769
12.0994 | -105.1757
12.1031 | -105.1771
12.1069 | -105.173
12.1106 | -105.1683
12.1144 | -105.1675
12.1181 | -105.1588
12.1219 | -105.1496
12.1256 | -105.1435
12.1294 | -105.1355
12.1331 | -105.1266
12.1369 | -105.1197
12.1406 | -105.1121
12.1444 | -105.1058
12.1481 | -105.0974
12.1519 | -105.0912
12.1556 | -105.0877
12.1594 | -105.0848
12.1631 | -105.0782
12.1669 | -105.0725
12.1707 | -105.0706
12.1744 | -105.0633
12.1782 | -105.059
12.1819 | -105.0569
12.1857 | -105.0544
12.1894 | -105.0498
12.1932 | -105.0455
12.1969 | -105.0451
12.2007 | -105.0428
12.2044 | -105.0379
12.2082 | -105.0375
12.2119 | -105.0375
12.2157 | -105.0364
12.2194 | -105.035
12.2232 | -105.0356
12.2269 | -105.0376
12.2307 | -105.0405
12.2344 | -105.0438
12.2382 | -105.0508
12.2419 | -105.0558
12.2457 | -105.0597
12.2494 | -105.0697
12.2532 | -105.0801
12.2569 | -105.0931
12.2607 | -105.1068
12.2644 | -105.1216
12.2682 | -105.1331
12.2719 | -105.1465
12.2757 | -105.1613
12.2794 | -105.1726
12.2832 | -105.1854
12.2869 | -105.1972
12.2907 | -105.2038
12.2944 | -105.2063
12.2982 | -105.2086
12.3019 | -105.2059
12.3057 | -105.2027
12.3094 | -105.2003
12.3132 | -105.1923
12.3169 | -105.1829
12.3207 | -105.1691
12.3244 | -105.1554
12.3282 | -105.1418
12.3319 | -105.1281
12.3357 | -105.1139
12.3394 | -105.1042
12.3432 | -105.0903
12.3469 | -105.0739
12.3507 | -105.0608
12.3544 | -105.0495
12.3582 | -105.0401
12.3619 | -105.0316
12.3657 | -105.0297
12.3694 | -105.0309
12.3732 | -105.0325
12.3769 | -105.0359
12.3807 | -105.0418
12.3844 | -105.0467
12.3882 | -105.0538
12.3919 | -105.0593
12.3957 | -105.0656
12.3994 | -105.0708
12.4032 | -105.0769
12.4069 | -105.084
12.4107 | -105.0914
12.4144 | -105.0983
12.4182 | -105.1072
12.4219 | -105.1129
12.4257 | -105.1147
12.4294 | -105.114
12.4332 | -105.1114
12.4369 | -105.106
12.4407 | -105.1012
12.4444 | -105.0949
12.4482 | -105.0945
12.4519 | -105.0912
12.4557 | -105.0851
12.4594 | -105.0788
12.4632 | -105.0706
12.4669 | -105.0611
12.4707 | -105.0556
12.4744 | -105.0528
12.4782 | -105.0445
12.4819 | -105.034
12.4857 | -105.025
12.4895 | -105.015
12.4932 | -105.0081
12.497 | -105.0014
12.5007 | -104.9952
12.5045 | -104.9903
12.5082 | -104.9859
12.512 | -104.9824
12.5157 | -104.9788
12.5195 | -104.9781
12.5232 | -104.9814
12.527 | -104.9871
12.5307 | -104.9959
12.5345 | -105.0015
12.5382 | -105.0073
12.542 | -105.0194
12.5457 | -105.0309
12.5495 | -105.042
12.5532 | -105.0558
12.557 | -105.0694
12.5607 | -105.0798
12.5645 | -105.0895
12.5682 | -105.0985
12.572 | -105.1095
12.5757 | -105.1155
12.5795 | -105.1214
12.5832 | -105.1245
12.587 | -105.1255
12.5907 | -105.1243
12.5945 | -105.1195
12.5982 | -105.1124
12.602 | -105.1047
12.6057 | -105.0937
12.6095 | -105.0826
12.6132 | -105.0675
12.617 | -105.0497
12.6207 | -105.0305
12.6245 | -105.0077
12.6282 | -104.9852
12.632 | -104.9602
12.6357 | -104.936
12.6395 | -104.9136
12.6432 | -104.8908
12.647 | -104.8671
12.6507 | -104.846
12.6545 | -104.823
12.6582 | -104.8017
12.662 | -104.7833
12.6657 | -104.7696
12.6695 | -104.7569
12.6732 | -104.7428
12.677 | -104.7332
12.6807 | -104.7281
12.6845 | -104.7258
12.6882 | -104.7236
12.692 | -104.724
12.6957 | -104.7282
12.6995 | -104.7359
12.7032 | -104.7436
12.707 | -104.7542
12.7107 | -104.7666
12.7145 | -104.7788
12.7182 | -104.7965
12.722 | -104.8169
12.7257 | -104.833
12.7295 | -104.851
12.7332 | -104.8674
12.737 | -104.8849
12.7407 | -104.9083
12.7445 | -104.9335
12.7482 | -104.957
12.752 | -104.9799
12.7557 | -105.0017
12.7595 | -105.0231
12.7632 | -105.0394
12.767 | -105.0548
12.7707 | -105.0711
12.7745 | -105.0872
12.7782 | -105.1007
12.782 | -105.1139
12.7857 | -105.126
12.7895 | -105.1362
12.7932 | -105.1455
12.797 | -105.1576
12.8008 | -105.1682
12.8045 | -105.1787
12.8083 | -105.1889
12.812 | -105.1985
12.8158 | -105.2087
12.8195 | -105.2186
12.8233 | -105.2291
12.827 | -105.2424
12.8308 | -105.2559
12.8345 | -105.2679
12.8383 | -105.2815
12.842 | -105.2919
12.8458 | -105.3039
12.8495 | -105.3166
12.8533 | -105.3266
12.857 | -105.34
12.8608 | -105.3479
12.8645 | -105.3497
12.8683 | -105.3566
12.872 | -105.3659
12.8758 | -105.3757
12.8795 | -105.3833
12.8833 | -105.3935
12.887 | -105.4063
12.8908 | -105.4192
12.8945 | -105.4323
12.8983 | -105.449
12.902 | -105.4644
12.9058 | -105.4789
12.9095 | -105.4947
12.9133 | -105.5126
12.917 | -105.5354
12.9208 | -105.5565
12.9245 | -105.5722
12.9283 | -105.5911
12.932 | -105.6087
12.9358 | -105.6243
12.9395 | -105.6376
12.9433 | -105.6522
12.947 | -105.6686
12.9508 | -105.6803
12.9545 | -105.6898
12.9583 | -105.6969
12.962 | -105.7053
12.9658 | -105.7088
12.9695 | -105.7092
12.9733 | -105.7072
12.977 | -105.7026
12.9808 | -105.6953
12.9845 | -105.6824
12.9883 | -105.6762
12.992 | -105.664
12.9958 | -105.6528
12.9995 | -105.6383
13.0033 | -105.6274
13.007 | -105.6166
13.0108 | -105.6023
13.0145 | -105.5841
13.0183 | -105.5698
13.022 | -105.5538
13.0258 | -105.5386
13.0295 | -105.5251
13.0333 | -105.5104
13.037 | -105.4958
13.0408 | -105.4834
13.0445 | -105.4713
13.0483 | -105.459
13.052 | -105.4491
13.0558 | -105.438
13.0595 | -105.4254
13.0633 | -105.4125
13.067 | -105.3988
13.0708 | -105.3867
13.0745 | -105.3734
13.0783 | -105.3606
13.082 | -105.3468
13.0858 | -105.3329
13.0895 | -105.3189
13.0933 | -105.3081
13.097 | -105.2981
13.1008 | -105.2896
13.1045 | -105.2793
13.1083 | -105.27
13.112 | -105.261
13.1158 | -105.2501
13.1196 | -105.2401
13.1233 | -105.231
13.1271 | -105.2202
13.1308 | -105.2097
13.1346 | -105.2023
13.1383 | -105.1945
13.1421 | -105.1831
13.1458 | -105.1753
13.1496 | -105.168
13.1533 | -105.1617
13.1571 | -105.1554
13.1608 | -105.1494
13.1646 | -105.1434
13.1683 | -105.1398
13.1721 | -105.1341
13.1758 | -105.1289
13.1796 | -105.1242
13.1833 | -105.1186
13.1871 | -105.112
13.1908 | -105.1082
13.1946 | -105.1045
13.1983 | -105.102
13.2021 | -105.1037
13.2058 | -105.109
13.2096 | -105.1099
13.2133 | -105.1135
13.2171 | -105.1156
13.2208 | -105.1185
13.2246 | -105.1235
13.2283 | -105.1305
13.2321 | -105.1365
13.2358 | -105.1429
13.2396 | -105.1528
13.2433 | -105.1611
13.2471 | -105.164
13.2508 | -105.1695
13.2546 | -105.1753
13.2583 | -105.1843
13.2621 | -105.1863
13.2658 | -105.187
13.2696 | -105.186
13.2733 | -105.1843
13.2771 | -105.1805
13.2808 | -105.1763
13.2846 | -105.168
13.2883 | -105.1569
13.2921 | -105.1377
13.2958 | -105.1195
13.2996 | -105.1015
13.3033 | -105.0817
13.3071 | -105.0591
13.3108 | -105.0348
13.3146 | -105.0068
13.3183 | -104.9762
13.3221 | -104.9454
13.3258 | -104.9095
13.3296 | -104.8731
13.3333 | -104.8364
13.3371 | -104.7967
13.3408 | -104.7602
13.3446 | -104.7215
13.3483 | -104.6812
13.3521 | -104.6399
13.3558 | -104.5977
13.3596 | -104.5533
13.3633 | -104.5068
13.3671 | -104.4612
13.3708 | -104.4192
13.3746 | -104.3772
13.3783 | -104.3305
13.3821 | -104.2863
13.3858 | -104.2433
13.3896 | -104.1994
13.3933 | -104.1557
13.3971 | -104.1159
13.4008 | -104.0789
13.4046 | -104.0432
13.4083 | -104.011
13.4121 | -103.9843
13.4158 | -103.9622
13.4196 | -103.9432
13.4233 | -103.9299
13.4271 | -103.9218
13.4308 | -103.9188
13.4346 | -103.9224
13.4383 | -103.933
13.4421 | -103.9485
13.4459 | -103.9693
13.4496 | -103.9933
13.4534 | -104.0264
13.4571 | -104.0659
13.4609 | -104.1071
13.4646 | -104.1511
13.4684 | -104.2014
13.4721 | -104.258
13.4759 | -104.3186
13.4796 | -104.3834
13.4834 | -104.4531
13.4871 | -104.523
13.4909 | -104.593
13.4946 | -104.6653
13.4984 | -104.7341
13.5021 | -104.8034
13.5059 | -104.8746
13.5096 | -104.9427
13.5134 | -105.0028
13.5171 | -105.0594
13.5209 | -105.1119
13.5246 | -105.1294
13.5284 | -105.1736
13.5321 | -105.2093
13.5359 | -105.2332
13.5396 | -105.2582
13.5434 | -105.2802
13.5471 | -105.2971
13.5509 | -105.3079
13.5546 | -105.3159
13.5584 | -105.3226
13.5621 | -105.3184
13.5659 | -105.3105
13.5696 | -105.3039
13.5734 | -105.298
13.5771 | -105.2873
13.5809 | -105.2747
13.5846 | -105.2622
13.5884 | -105.2521
13.5921 | -105.2417
13.5959 | -105.2331
13.5996 | -105.2272
13.6034 | -105.2203
13.6071 | -105.2131
13.6109 | -105.2059
13.6146 | -105.2004
13.6184 | -105.1943
13.6221 | -105.1883
13.6259 | -105.1849
13.6296 | -105.1829
13.6334 | -105.1817
13.6371 | -105.1789
13.6409 | -105.1772
13.6446 | -105.1761
13.6484 | -105.1759
13.6521 | -105.173
13.6559 | -105.1709
13.6596 | -105.1707
13.6634 | -105.1668
13.6671 | -105.1612
13.6709 | -105.1569
13.6746 | -105.1532
13.6784 | -105.1501
13.6821 | -105.1501
13.6859 | -105.1497
13.6896 | -105.1523
13.6934 | -105.1574
13.6971 | -105.1609
13.7009 | -105.1647
13.7046 | -105.1713
13.7084 | -105.1829
13.7121 | -105.1929
13.7159 | -105.2026
13.7196 | -105.2145
13.7234 | -105.2279
13.7271 | -105.2407
13.7309 | -105.2572
13.7346 | -105.2789
13.7384 | -105.297
13.7421 | -105.3113
13.7459 | -105.3239
13.7496 | -105.3401
13.7534 | -105.357
13.7571 | -105.3694
13.7609 | -105.3832
13.7647 | -105.3969
13.7684 | -105.4092
13.7722 | -105.4216
13.7759 | -105.4277
13.7797 | -105.4315
13.7834 | -105.4367
13.7872 | -105.4398
13.7909 | -105.441
13.7947 | -105.4416
13.7984 | -105.4397
13.8022 | -105.4385
13.8059 | -105.4331
13.8097 | -105.4272
13.8134 | -105.4232
13.8172 | -105.4185
13.8209 | -105.4118
13.8247 | -105.4055
13.8284 | -105.4005
13.8322 | -105.395
13.8359 | -105.3879
13.8397 | -105.3784
13.8434 | -105.3737
13.8472 | -105.3686
13.8509 | -105.3669
13.8547 | -105.3669
13.8584 | -105.3649
13.8622 | -105.3643
13.8659 | -105.3648
13.8697 | -105.367
13.8734 | -105.3703
13.8772 | -105.3749
13.8809 | -105.3781
13.8847 | -105.3833
13.8884 | -105.388
13.8922 | -105.3936
13.8959 | -105.4017
13.8997 | -105.4104
13.9034 | -105.4208
13.9072 | -105.4289
13.9109 | -105.4405
13.9147 | -105.4513
13.9184 | -105.4633
13.9222 | -105.4741
13.9259 | -105.4842
13.9297 | -105.4952
13.9334 | -105.5049
13.9372 | -105.5167
13.9409 | -105.5277
13.9447 | -105.5393
13.9484 | -105.5482
13.9522 | -105.5526
13.9559 | -105.5619
13.9597 | -105.5655
13.9634 | -105.5552
13.9672 | -105.5389
13.9709 | -105.5285
13.9747 | -105.5235
13.9784 | -105.5135
13.9822 | -105.501
13.9859 | -105.4836
13.9897 | -105.4612
13.9934 | -105.4376
13.9972 | -105.4143
14.0009 | -105.3878
14.0047 | -105.3597
14.0084 | -105.3301
14.0122 | -105.2965
14.0159 | -105.2622
14.0197 | -105.2305
14.0234 | -105.1949
14.0272 | -105.1607
14.0309 | -105.1302
14.0347 | -105.0998
14.0384 | -105.0717
14.0422 | -105.0478
14.0459 | -105.0276
14.0497 | -105.0105
14.0534 | -104.9969
14.0572 | -104.9865
14.0609 | -104.9806
14.0647 | -104.9956
14.0684 | -104.9958
14.0722 | -104.9993
14.0759 | -105.007
14.0797 | -105.0169
14.0835 | -105.0308
14.0872 | -105.0466
14.091 | -105.0655
14.0947 | -105.0869
14.0985 | -105.105
14.1022 | -105.1273
14.106 | -105.1504
14.1097 | -105.1745
14.1135 | -105.1983
14.1172 | -105.224
14.121 | -105.2453
14.1247 | -105.2637
14.1285 | -105.2794
14.1322 | -105.2939
14.136 | -105.31
14.1397 | -105.3258
14.1435 | -105.3437
14.1472 | -105.3549
14.151 | -105.3661
14.1547 | -105.3785
14.1585 | -105.3875
14.1622 | -105.3977
14.166 | -105.4093
14.1697 | -105.4199
14.1735 | -105.4194
14.1772 | -105.4151
14.181 | -105.4123
14.1847 | -105.4081
14.1885 | -105.4055
14.1922 | -105.4025
14.196 | -105.3978
14.1997 | -105.39
14.2035 | -105.3834
14.2072 | -105.3792
14.211 | -105.375
14.2147 | -105.3686
14.2185 | -105.3656
14.2222 | -105.3632
14.226 | -105.3608
14.2297 | -105.3561
14.2335 | -105.3533
14.2372 | -105.3546
14.241 | -105.3583
14.2447 | -105.3568
14.2485 | -105.3554
14.2522 | -105.3554
14.256 | -105.36
14.2597 | -105.3679
14.2635 | -105.3756
14.2672 | -105.3851
14.271 | -105.3929
14.2747 | -105.4017
14.2785 | -105.4107
14.2822 | -105.4196
14.286 | -105.4291
14.2897 | -105.4409
14.2935 | -105.4521
14.2972 | -105.4608
14.301 | -105.4724
14.3047 | -105.4867
14.3085 | -105.495
14.3122 | -105.5021
14.316 | -105.5122
14.3197 | -105.5213
14.3235 | -105.5288
14.3272 | -105.5386
14.331 | -105.5486
14.3347 | -105.5611
14.3385 | -105.5703
14.3422 | -105.5763
14.346 | -105.5797
14.3497 | -105.5821
14.3535 | -105.5892
14.3572 | -105.5896
14.361 | -105.59
14.3647 | -105.5859
14.3685 | -105.5812
14.3722 | -105.5791
14.376 | -105.5766
14.3797 | -105.5727
14.3835 | -105.567
14.3872 | -105.5607
14.391 | -105.5513
14.3947 | -105.5416
14.3985 | -105.5324
14.4023 | -105.5185
14.406 | -105.5034
14.4098 | -105.4895
14.4135 | -105.4747
14.4173 | -105.4571
14.421 | -105.4393
14.4248 | -105.4202
14.4285 | -105.4033
14.4323 | -105.3892
14.436 | -105.3745
14.4398 | -105.3591
14.4435 | -105.3439
14.4473 | -105.3279
14.451 | -105.312
14.4548 | -105.2983
14.4585 | -105.2848
14.4623 | -105.2708
14.466 | -105.2597
14.4698 | -105.2488
14.4735 | -105.2415
14.4773 | -105.2376
14.481 | -105.2329
14.4848 | -105.23
14.4885 | -105.2301
14.4923 | -105.2316
14.496 | -105.2302
14.4998 | -105.2314
14.5035 | -105.2356
14.5073 | -105.2393
14.511 | -105.2442
14.5148 | -105.2514
14.5185 | -105.2612
14.5223 | -105.2697
14.526 | -105.2789
14.5298 | -105.2904
14.5335 | -105.302
14.5373 | -105.3146
14.541 | -105.3253
14.5448 | -105.3339
14.5485 | -105.3442
14.5523 | -105.3541
14.556 | -105.3638
14.5598 | -105.3731
14.5635 | -105.3813
14.5673 | -105.3856
14.571 | -105.3925
14.5748 | -105.3953
14.5785 | -105.3957
14.5823 | -105.3959
14.586 | -105.3999
14.5898 | -105.4
14.5935 | -105.3965
14.5973 | -105.3907
14.601 | -105.3848
14.6048 | -105.3809
14.6085 | -105.3708
14.6123 | -105.3661
14.616 | -105.3591
14.6198 | -105.3473
14.6235 | -105.3318
14.6273 | -105.3204
14.631 | -105.3102
14.6348 | -105.2948
14.6385 | -105.2767
14.6423 | -105.2597
14.646 | -105.2435
14.6498 | -105.2247
14.6535 | -105.2058
14.6573 | -105.1861
14.661 | -105.1658
14.6648 | -105.1435
14.6685 | -105.1219
14.6723 | -105.0979
14.676 | -105.0718
14.6798 | -105.0493
14.6835 | -105.0268
14.6873 | -105.0034
14.691 | -104.9863
14.6948 | -104.9699
14.6985 | -104.9543
14.7023 | -104.9416
14.706 | -104.9329
14.7098 | -104.9277
14.7135 | -104.9252
14.7173 | -104.9241
14.721 | -104.9254
14.7248 | -104.9262
14.7286 | -104.9309
14.7323 | -104.9387
14.7361 | -104.9511
14.7398 | -104.9643
14.7436 | -104.9797
14.7473 | -104.9955
14.7511 | -105.0156
14.7548 | -105.0402
14.7586 | -105.0625
14.7623 | -105.0869
14.7661 | -105.1125
14.7698 | -105.135
14.7736 | -105.1595
14.7773 | -105.1843
14.7811 | -105.2062
14.7848 | -105.2266
14.7886 | -105.244
14.7923 | -105.2579
14.7961 | -105.2671
14.7998 | -105.2791
14.8036 | -105.2885
14.8073 | -105.2994
14.8111 | -105.3065
14.8148 | -105.3069
14.8186 | -105.3056
14.8223 | -105.3025
14.8261 | -105.2974
14.8298 | -105.2913
14.8336 | -105.2813
14.8373 | -105.2709
14.8411 | -105.2585
14.8448 | -105.2439
14.8486 | -105.2313
14.8523 | -105.217
14.8561 | -105.1985
14.8598 | -105.1812
14.8636 | -105.1688
14.8673 | -105.1542
14.8711 | -105.1386
14.8748 | -105.1277
14.8786 | -105.1176
14.8823 | -105.1066
14.8861 | -105.0989
14.8898 | -105.0921
14.8936 | -105.086
14.8973 | -105.0804
14.9011 | -105.0768
14.9048 | -105.0717
14.9086 | -105.0682
14.9123 | -105.0674
14.9161 | -105.0687
14.9198 | -105.0716
14.9236 | -105.0737
14.9273 | -105.0757
14.9311 | -105.078
14.9348 | -105.0805
14.9386 | -105.0866
14.9423 | -105.0899
14.9461 | -105.091
14.9498 | -105.089
14.9536 | -105.0878
14.9573 | -105.0838
14.9611 | -105.0807
14.9648 | -105.0757
14.9686 | -105.07
14.9723 | -105.0648
14.9761 | -105.0561
14.9798 | -105.0469
14.9836 | -105.0387
14.9873 | -105.0302
14.9911 | -105.0196
14.9948 | -105.0114
14.9986 | -105.0045
15.0023 | -104.9943
15.0061 | -104.9822
15.0098 | -104.9718
15.0136 | -104.9583
15.0173 | -104.9464
15.0211 | -104.9363
15.0248 | -104.9255
15.0286 | -104.9167
15.0323 | -104.9082
15.0361 | -104.8985
15.0398 | -104.8886
15.0436 | -104.8802
15.0474 | -104.8748
15.0511 | -104.8699
15.0549 | -104.8626
15.0586 | -104.8571
15.0624 | -104.8509
15.0661 | -104.8485
15.0699 | -104.8498
15.0736 | -104.8528
15.0774 | -104.8565
15.0811 | -104.8631
15.0849 | -104.8705
15.0886 | -104.8781
15.0924 | -104.8869
15.0961 | -104.897
15.0999 | -104.9073
15.1036 | -104.9206
15.1074 | -104.9344
15.1111 | -104.9494
15.1149 | -104.9666
15.1186 | -104.9848
15.1224 | -105.0043
15.1261 | -105.0227
15.1299 | -105.0388
15.1336 | -105.0535
15.1374 | -105.0694
15.1411 | -105.0849
15.1449 | -105.0987
15.1486 | -105.0959
15.1524 | -105.1031
15.1561 | -105.1113
15.1599 | -105.1177
15.1636 | -105.1228
15.1674 | -105.1263
15.1711 | -105.126
15.1749 | -105.1248
15.1786 | -105.1225
15.1824 | -105.117
15.1861 | -105.1055
15.1899 | -105.0945
15.1936 | -105.0817
15.1974 | -105.0693
15.2011 | -105.0556
15.2049 | -105.0354
15.2086 | -105.0171
15.2124 | -105.0004
15.2161 | -104.9815
15.2199 | -104.9638
15.2236 | -104.947
15.2274 | -104.9277
15.2311 | -104.9087
15.2349 | -104.8893
15.2386 | -104.8687
15.2424 | -104.8507
15.2461 | -104.8305
15.2499 | -104.8127
15.2536 | -104.8004
15.2574 | -104.7882
15.2611 | -104.7784
15.2649 | -104.7686
15.2686 | -104.7615
15.2724 | -104.7565
15.2761 | -104.752
15.2799 | -104.7525
15.2836 | -104.7502
15.2874 | -104.7484
15.2911 | -104.7492
15.2949 | -104.7494
15.2986 | -104.7492
15.3024 | -104.7522
15.3061 | -104.7559
15.3099 | -104.7605
15.3136 | -104.7667
15.3174 | -104.7731
15.3211 | -104.7769
15.3249 | -104.7783
15.3286 | -104.7797
15.3324 | -104.7807
15.3361 | -104.7811
15.3399 | -104.7804
15.3436 | -104.7798
15.3474 | -104.7775
15.3511 | -104.7705
15.3549 | -104.7626
15.3587 | -104.7488
15.3624 | -104.7307
15.3662 | -104.7174
15.3699 | -104.7007
15.3737 | -104.6797
15.3774 | -104.6522
15.3812 | -104.624
15.3849 | -104.5969
15.3887 | -104.5701
15.3924 | -104.542
15.3962 | -104.5125
15.3999 | -104.4806
15.4037 | -104.4489
15.4074 | -104.4196
15.4112 | -104.3884
15.4149 | -104.3589
15.4187 | -104.3263
15.4224 | -104.293
15.4262 | -104.2646
15.4299 | -104.2365
15.4337 | -104.2087
15.4374 | -104.1845
15.4412 | -104.1628
15.4449 | -104.1452
15.4487 | -104.1275
15.4524 | -104.1118
15.4562 | -104.1018
15.4599 | -104.092
15.4637 | -104.0834
15.4674 | -104.0816
15.4712 | -104.0807
15.4749 | -104.0772
15.4787 | -104.0767
15.4824 | -104.0794
15.4862 | -104.0847
15.4899 | -104.0919
15.4937 | -104.1013
15.4974 | -104.1125
15.5012 | -104.1243
15.5049 | -104.1367
15.5087 | -104.1512
15.5124 | -104.1643
15.5162 | -104.1783
15.5199 | -104.1936
15.5237 | -104.2069
15.5274 | -104.2189
15.5312 | -104.2318
15.5349 | -104.2447
15.5387 | -104.2577
15.5424 | -104.2705
15.5462 | -104.2835
15.5499 | -104.2964
15.5537 | -104.3098
15.5574 | -104.3226
15.5612 | -104.334
15.5649 | -104.3467
15.5687 | -104.3614
15.5724 | -104.3806
15.5762 | -104.3969
15.5799 | -104.4145
15.5837 | -104.4351
15.5874 | -104.4512
15.5912 | -104.4714
15.5949 | -104.4883
15.5987 | -104.5062
15.6024 | -104.5244
15.6062 | -104.5423
15.6099 | -104.5587
15.6137 | -104.5747
15.6174 | -104.5917
15.6212 | -104.6077
15.6249 | -104.6216
15.6287 | -104.6344
15.6324 | -104.6438
15.6362 | -104.6512
15.6399 | -104.6596
15.6437 | -104.6651
15.6474 | -104.6638
15.6512 | -104.661
15.6549 | -104.6547
15.6587 | -104.6453
15.6624 | -104.6317
15.6662 | -104.6118
15.6699 | -104.5949
15.6737 | -104.5739
15.6775 | -104.5466
15.6812 | -104.5161
15.685 | -104.4838
15.6887 | -104.4384
15.6925 | -104.4027
15.6962 | -104.3674
15.7 | -104.3313
15.7037 | -104.2935
15.7075 | -104.2537
15.7112 | -104.214
15.715 | -104.1721
15.7187 | -104.1342
15.7225 | -104.0964
15.7262 | -104.0564
15.73 | -104.0201
15.7337 | -103.9852
15.7375 | -103.9535
15.7412 | -103.921
15.745 | -103.8928
15.7487 | -103.8694
15.7525 | -103.8486
15.7562 | -103.8306
15.76 | -103.8153
15.7637 | -103.7992
15.7675 | -103.7817
15.7712 | -103.7684
15.775 | -103.7595
15.7787 | -103.7508
15.7825 | -103.7455
15.7862 | -103.7438
15.79 | -103.7437
15.7937 | -103.7453
15.7975 | -103.7459
15.8012 | -103.75
15.805 | -103.7537
15.8087 | -103.7555
15.8125 | -103.758
15.8162 | -103.7587
15.82 | -103.7635
15.8237 | -103.7673
15.8275 | -103.7708
15.8312 | -103.7724
15.835 | -103.7755
15.8387 | -103.7794
15.8425 | -103.7815
15.8462 | -103.7827
15.85 | -103.7878
15.8537 | -103.7921
15.8575 | -103.7964
15.8612 | -103.8013
15.865 | -103.8025
15.8687 | -103.8025
15.8725 | -103.8061
15.8762 | -103.808
15.88 | -103.8092
15.8837 | -103.8123
15.8875 | -103.8186
15.8912 | -103.8257
15.895 | -103.831
15.8987 | -103.8378
15.9025 | -103.8449
15.9062 | -103.8543
15.91 | -103.8655
15.9137 | -103.873
15.9175 | -103.8796
15.9212 | -103.8893
15.925 | -103.8974
15.9287 | -103.9048
15.9325 | -103.9112
15.9362 | -103.9169
15.94 | -103.9234
15.9437 | -103.9318
15.9475 | -103.9346
15.9512 | -103.9373
15.955 | -103.9428
15.9587 | -103.9456
15.9625 | -103.9451
15.9662 | -103.941
15.97 | -103.9413
15.9737 | -103.9386
15.9775 | -103.9338
15.9812 | -103.9288
15.985 | -103.9238
15.9887 | -103.9194
15.9925 | -103.9172
15.9962 | -103.9172
16 | -103.9158
16.0038 | -103.912
16.0075 | -103.9104
16.0113 | -103.9083
16.015 | -103.9062
16.0188 | -103.9049
16.0225 | -103.9057
16.0263 | -103.9089
16.03 | -103.9119
16.0338 | -103.9177
16.0375 | -103.9224
16.0413 | -103.9257
16.045 | -103.9321
16.0488 | -103.9384
16.0525 | -103.9469
16.0563 | -103.9544
16.06 | -103.9625
16.0638 | -103.9695
16.0675 | -103.9744
16.0713 | -103.9792
16.075 | -103.9815
16.0788 | -103.9804
16.0825 | -103.9815
16.0863 | -103.9824
16.09 | -103.9817
16.0938 | -103.9822
16.0975 | -103.9837
16.1013 | -103.9796
16.105 | -103.9738
16.1088 | -103.966
16.1125 | -103.9576
16.1163 | -103.9486
16.12 | -103.9404
16.1238 | -103.9302
16.1275 | -103.9148
16.1313 | -103.8972
16.135 | -103.8796
16.1388 | -103.8608
16.1425 | -103.8414
16.1463 | -103.8192
16.15 | -103.7964
16.1538 | -103.7749
16.1575 | -103.7557
16.1613 | -103.737
16.165 | -103.7169
16.1688 | -103.6949
16.1725 | -103.6718
16.1763 | -103.6508
16.18 | -103.6331
16.1838 | -103.6177
16.1875 | -103.6047
16.1913 | -103.5933
16.195 | -103.586
16.1988 | -103.5781
16.2025 | -103.5731
16.2063 | -103.5703
16.21 | -103.5686
16.2138 | -103.5689
16.2175 | -103.5733
16.2213 | -103.5782
16.225 | -103.5885
16.2288 | -103.5996
16.2325 | -103.6187
16.2363 | -103.6315
16.24 | -103.6434
16.2438 | -103.6594
16.2475 | -103.6766
16.2513 | -103.6946
16.255 | -103.7141
16.2588 | -103.7367
16.2625 | -103.7588
16.2663 | -103.7815
16.27 | -103.8059
16.2738 | -103.831
16.2775 | -103.8525
16.2813 | -103.8754
16.285 | -103.8979
16.2888 | -103.9204
16.2925 | -103.9433
16.2963 | -103.9669
16.3 | -103.9919
16.3038 | -104.0157
16.3075 | -104.0411
16.3113 | -104.0646
16.315 | -104.0848
16.3188 | -104.104
16.3225 | -104.1211
16.3263 | -104.1358
16.3301 | -104.1467
16.3338 | -104.1579
16.3376 | -104.1682
16.3413 | -104.1784
16.3451 | -104.1837
16.3488 | -104.1876
16.3526 | -104.1905
16.3563 | -104.1931
16.3601 | -104.1994
16.3638 | -104.2023
16.3676 | -104.1989
16.3713 | -104.1971
16.3751 | -104.1984
16.3788 | -104.2002
16.3826 | -104.1987
16.3863 | -104.1952
16.3901 | -104.1908
16.3938 | -104.1851
16.3976 | -104.1821
16.4013 | -104.1797
16.4051 | -104.1771
16.4088 | -104.1772
16.4126 | -104.1765
16.4163 | -104.1734
16.4201 | -104.1711
16.4238 | -104.1725
16.4276 | -104.1746
16.4313 | -104.1759
16.4351 | -104.177
16.4388 | -104.1785
16.4426 | -104.1835
16.4463 | -104.1888
16.4501 | -104.1927
16.4538 | -104.2011
16.4576 | -104.2086
16.4613 | -104.2174
16.4651 | -104.228
16.4688 | -104.2407
16.4726 | -104.2533
16.4763 | -104.2686
16.4801 | -104.281
16.4838 | -104.296
16.4876 | -104.3111
16.4913 | -104.3278
16.4951 | -104.3452
16.4988 | -104.3662
16.5026 | -104.3879
16.5063 | -104.4128
16.5101 | -104.4394
16.5138 | -104.4687
16.5176 | -104.498
16.5213 | -104.5281
16.5251 | -104.5591
16.5288 | -104.5913
16.5326 | -104.6236
16.5363 | -104.6571
16.5401 | -104.6916
16.5438 | -104.7257
16.5476 | -104.7608
16.5513 | -104.7967
16.5551 | -104.8329
16.5588 | -104.8663
16.5626 | -104.8985
16.5663 | -104.9333
16.5701 | -104.9668
16.5738 | -104.9951
16.5776 | -105.0215
16.5813 | -105.0479
16.5851 | -105.0685
16.5888 | -105.0879
16.5926 | -105.1062
16.5963 | -105.1185
16.6001 | -105.1302
16.6038 | -105.1419
16.6076 | -105.1477
16.6113 | -105.1469
16.6151 | -105.1451
16.6188 | -105.143
16.6226 | -105.1375
16.6263 | -105.1278
16.6301 | -105.1166
16.6338 | -105.1012
16.6376 | -105.0847
16.6413 | -105.0652
16.6451 | -105.0408
16.6489 | -105.0123
16.6526 | -104.9813
16.6564 | -104.9531
16.6601 | -104.9201
16.6639 | -104.8876
16.6676 | -104.8545
16.6714 | -104.821
16.6751 | -104.7869
16.6789 | -104.7514
16.6826 | -104.7183
16.6864 | -104.6865
16.6901 | -104.6571
16.6939 | -104.6289
16.6976 | -104.6004
16.7014 | -104.5722
16.7051 | -104.5433
16.7089 | -104.5131
16.7126 | -104.4795
16.7164 | -104.4528
16.7201 | -104.4296
16.7239 | -104.4081
16.7276 | -104.3878
16.7314 | -104.3679
16.7351 | -104.35
16.7389 | -104.3343
16.7426 | -104.3185
16.7464 | -104.3044
16.7501 | -104.29
16.7539 | -104.2752
16.7576 | -104.2615
16.7614 | -104.249
16.7651 | -104.2374
16.7689 | -104.2257
16.7726 | -104.2216
16.7764 | -104.2136
16.7801 | -104.2065
16.7839 | -104.2013
16.7876 | -104.1971
16.7914 | -104.1937
16.7951 | -104.1927
16.7989 | -104.1958
16.8026 | -104.1998
16.8064 | -104.2036
16.8101 | -104.2074
16.8139 | -104.2115
16.8176 | -104.2178
16.8214 | -104.2274
16.8251 | -104.241
16.8289 | -104.2541
16.8326 | -104.2697
16.8364 | -104.2863
16.8401 | -104.3045
16.8439 | -104.3229
16.8476 | -104.3388
16.8514 | -104.355
16.8551 | -104.3704
16.8589 | -104.3849
16.8626 | -104.4006
16.8664 | -104.4176
16.8701 | -104.4331
16.8739 | -104.4456
16.8776 | -104.4563
16.8814 | -104.4673
16.8851 | -104.4736
16.8889 | -104.4802
16.8926 | -104.4882
16.8964 | -104.4908
16.9001 | -104.491
16.9039 | -104.4883
16.9076 | -104.4859
16.9114 | -104.4815
16.9151 | -104.4736
16.9189 | -104.4625
16.9226 | -104.4477
16.9264 | -104.432
16.9301 | -104.4155
16.9339 | -104.3976
16.9376 | -104.3809
16.9414 | -104.3665
16.9451 | -104.3526
16.9489 | -104.3412
16.9526 | -104.3326
16.9564 | -104.319
16.9601 | -104.3031
16.9639 | -104.287
16.9677 | -104.2714
16.9714 | -104.2587
16.9752 | -104.2465
16.9789 | -104.2353
16.9827 | -104.2256
16.9864 | -104.2144
16.9902 | -104.2067
16.9939 | -104.2017
16.9977 | -104.1975
17.0014 | -104.1947
17.0052 | -104.1898
17.0089 | -104.1882
17.0127 | -104.1878
17.0164 | -104.1889
17.0202 | -104.19
17.0239 | -104.1934
17.0277 | -104.1982
17.0314 | -104.2038
17.0352 | -104.2124
17.0389 | -104.2213
17.0427 | -104.2312
17.0464 | -104.2421
17.0502 | -104.2545
17.0539 | -104.2688
17.0577 | -104.2834
17.0614 | -104.296
17.0652 | -104.3061
17.0689 | -104.3155
17.0727 | -104.3234
17.0764 | -104.3291
17.0802 | -104.3361
17.0839 | -104.342
17.0877 | -104.3459
17.0914 | -104.3498
17.0952 | -104.3538
17.0989 | -104.3556
17.1027 | -104.3531
17.1064 | -104.3487
17.1102 | -104.3419
17.1139 | -104.3325
17.1177 | -104.3237
17.1214 | -104.3106
17.1252 | -104.2939
17.1289 | -104.278
17.1327 | -104.2587
17.1364 | -104.2356
17.1402 | -104.2138
17.1439 | -104.1939
17.1477 | -104.1705
17.1514 | -104.1461
17.1552 | -104.1232
17.1589 | -104.1012
17.1627 | -104.0791
17.1664 | -104.0556
17.1702 | -104.0327
17.1739 | -104.0064
17.1777 | -103.9787
17.1814 | -103.9517
17.1852 | -103.925
17.1889 | -103.901
17.1927 | -103.8771
17.1964 | -103.8552
17.2002 | -103.8349
17.2039 | -103.8157
17.2077 | -103.8015
17.2114 | -103.7904
17.2152 | -103.7798
17.2189 | -103.7685
17.2227 | -103.762
17.2264 | -103.7591
17.2302 | -103.7595
17.2339 | -103.7619
17.2377 | -103.768
17.2414 | -103.7778
17.2452 | -103.791
17.2489 | -103.8081
17.2527 | -103.826
17.2564 | -103.8444
17.2602 | -103.8625
17.2639 | -103.8856
17.2677 | -103.9086
17.2714 | -103.9312
17.2752 | -103.9553
17.2789 | -103.9782
17.2827 | -104.0005
17.2865 | -104.0229
17.2902 | -104.048
17.294 | -104.0707
17.2977 | -104.0894
17.3015 | -104.1101
17.3052 | -104.1319
17.309 | -104.1508
17.3127 | -104.1587
17.3165 | -104.1726
17.3202 | -104.1867
17.324 | -104.1974
17.3277 | -104.2031
17.3315 | -104.2073
17.3352 | -104.21
17.339 | -104.2113
17.3427 | -104.2122
17.3465 | -104.2142
17.3502 | -104.2164
17.354 | -104.2174
17.3577 | -104.2152
17.3615 | -104.2141
17.3652 | -104.2113
17.369 | -104.2083
17.3727 | -104.2047
17.3765 | -104.1987
17.3802 | -104.1907
17.384 | -104.1827
17.3877 | -104.1723
17.3915 | -104.1616
17.3952 | -104.1533
17.399 | -104.1432
17.4027 | -104.132
17.4065 | -104.1196
17.4102 | -104.1075
17.414 | -104.0985
17.4177 | -104.0891
17.4215 | -104.0828
17.4252 | -104.0797
17.429 | -104.0769
17.4327 | -104.0738
17.4365 | -104.0709
17.4402 | -104.0696
17.444 | -104.0686
17.4477 | -104.0699
17.4515 | -104.0708
17.4552 | -104.0727
17.459 | -104.0766
17.4627 | -104.0831
17.4665 | -104.0947
17.4702 | -104.1055
17.474 | -104.1153
17.4777 | -104.126
17.4815 | -104.1361
17.4852 | -104.1465
17.489 | -104.1577
17.4927 | -104.1673
17.4965 | -104.1755
17.5002 | -104.1854
17.504 | -104.1963
17.5077 | -104.2082
17.5115 | -104.2148
17.5152 | -104.2229
17.519 | -104.2337
17.5227 | -104.2425
17.5265 | -104.2499
17.5302 | -104.2571
17.534 | -104.2619
17.5377 | -104.2632
17.5415 | -104.2665
17.5452 | -104.2706
17.549 | -104.2718
17.5527 | -104.2702
17.5565 | -104.2717
17.5602 | -104.2741
17.564 | -104.2756
17.5677 | -104.2787
17.5715 | -104.2795
17.5752 | -104.2844
17.579 | -104.2911
17.5827 | -104.2954
17.5865 | -104.2976
17.5902 | -104.2976
17.594 | -104.297
17.5978 | -104.2943
17.6015 | -104.2917
17.6053 | -104.2913
17.609 | -104.2945
17.6128 | -104.2953
17.6165 | -104.2915
17.6203 | -104.2899
17.624 | -104.2915
17.6278 | -104.2933
17.6315 | -104.293
17.6353 | -104.2882
17.639 | -104.2824
17.6428 | -104.2768
17.6465 | -104.2693
17.6503 | -104.2644
17.654 | -104.2572
17.6578 | -104.2515
17.6615 | -104.2445
17.6653 | -104.2316
17.669 | -104.2166
17.6728 | -104.2012
17.6765 | -104.1848
17.6803 | -104.1675
17.684 | -104.1512
17.6878 | -104.1343
17.6915 | -104.1183
17.6953 | -104.1019
17.699 | -104.0836
17.7028 | -104.0677
17.7065 | -104.0514
17.7103 | -104.0324
17.714 | -104.015
17.7178 | -103.9978
17.7215 | -103.9818
17.7253 | -103.9681
17.729 | -103.9565
17.7328 | -103.9463
17.7365 | -103.9341
17.7403 | -103.9233
17.744 | -103.9155
17.7478 | -103.9095
17.7515 | -103.9039
17.7553 | -103.9003
17.759 | -103.897
17.7628 | -103.8958
17.7665 | -103.8965
17.7703 | -103.8976
17.774 | -103.899
17.7778 | -103.9018
17.7815 | -103.9032
17.7853 | -103.9073
17.789 | -103.9136
17.7928 | -103.92
17.7965 | -103.9272
17.8003 | -103.9329
17.804 | -103.9375
17.8078 | -103.9415
17.8115 | -103.945
17.8153 | -103.9492
17.819 | -103.9515
17.8228 | -103.9515
17.8265 | -103.9467
17.8303 | -103.9436
17.834 | -103.9409
17.8378 | -103.9366
17.8415 | -103.9333
17.8453 | -103.9283
17.849 | -103.9215
17.8528 | -103.9062
17.8565 | -103.8997
17.8603 | -103.8883
17.864 | -103.8717
17.8678 | -103.858
17.8715 | -103.8506
17.8753 | -103.8413
17.879 | -103.8287
17.8828 | -103.8142
17.8865 | -103.8009
17.8903 | -103.7913
17.894 | -103.7816
17.8978 | -103.7702
17.9015 | -103.7608
17.9053 | -103.7499
17.909 | -103.7401
17.9128 | -103.7291
17.9165 | -103.7175
17.9203 | -103.7099
17.9241 | -103.7039
17.9278 | -103.6985
17.9316 | -103.6943
17.9353 | -103.6918
17.9391 | -103.6885
17.9428 | -103.6844
17.9466 | -103.6795
17.9503 | -103.6753
17.9541 | -103.6731
17.9578 | -103.6724
17.9616 | -103.6709
17.9653 | -103.672
17.9691 | -103.6725
17.9728 | -103.6743
17.9766 | -103.6763
17.9803 | -103.6787
17.9841 | -103.6795
17.9878 | -103.6811
17.9916 | -103.6824
17.9953 | -103.6834
17.9991 | -103.6843
18.0028 | -103.6847
18.0066 | -103.6828
18.0103 | -103.6809
18.0141 | -103.6805
18.0178 | -103.6802
18.0216 | -103.6768
18.0253 | -103.6719
18.0291 | -103.6678
18.0328 | -103.6639
18.0366 | -103.6628
18.0403 | -103.6583
18.0441 | -103.6502
18.0478 | -103.6404
18.0516 | -103.6295
18.0553 | -103.6187
18.0591 | -103.6096
18.0628 | -103.5995
18.0666 | -103.5872
18.0703 | -103.5725
18.0741 | -103.5589
18.0778 | -103.5459
18.0816 | -103.5348
18.0853 | -103.5214
18.0891 | -103.5056
18.0928 | -103.4888
18.0966 | -103.4703
18.1003 | -103.4534
18.1041 | -103.4365
18.1078 | -103.4187
18.1116 | -103.4008
18.1153 | -103.3819
18.1191 | -103.3658
18.1228 | -103.3543
18.1266 | -103.3412
18.1303 | -103.3246
18.1341 | -103.3088
18.1378 | -103.2929
18.1416 | -103.2791
18.1453 | -103.2651
18.1491 | -103.2508
18.1528 | -103.2368
18.1566 | -103.2218
18.1603 | -103.2088
18.1641 | -103.1974
18.1678 | -103.1857
18.1716 | -103.1739
18.1753 | -103.1627
18.1791 | -103.1479
18.1828 | -103.1312
18.1866 | -103.1103
18.1903 | -103.0895
18.1941 | -103.0725
18.1978 | -103.0554
18.2016 | -103.0384
18.2053 | -103.0188
18.2091 | -102.9996
18.2128 | -102.9794
18.2166 | -102.9578
18.2203 | -102.9378
18.2241 | -102.9166
18.2278 | -102.895
18.2316 | -102.8729
18.2353 | -102.8523
18.2391 | -102.8331
18.2429 | -102.8171
18.2466 | -102.8023
18.2504 | -102.7855
18.2541 | -102.7685
18.2579 | -102.7535
18.2616 | -102.7379
18.2654 | -102.7207
18.2691 | -102.7021
18.2729 | -102.6837
18.2766 | -102.6665
18.2804 | -102.6511
18.2841 | -102.6361
18.2879 | -102.6218
18.2916 | -102.6053
18.2954 | -102.5869
18.2991 | -102.5712
18.3029 | -102.5558
18.3066 | -102.5396
18.3104 | -102.5234
18.3141 | -102.5072
18.3179 | -102.4884
18.3216 | -102.4682
18.3254 | -102.4512
18.3291 | -102.4345
18.3329 | -102.419
18.3366 | -102.4036
18.3404 | -102.3869
18.3441 | -102.3705
18.3479 | -102.3561
18.3516 | -102.3429
18.3554 | -102.3291
18.3591 | -102.3117
18.3629 | -102.2955
18.3666 | -102.2816
18.3704 | -102.264
18.3741 | -102.2452
18.3779 | -102.2263
18.3816 | -102.2091
18.3854 | -102.1934
18.3891 | -102.18
18.3929 | -102.165
18.3966 | -102.1529
18.4004 | -102.1382
18.4041 | -102.1252
18.4079 | -102.1138
18.4116 | -102.1001
18.4154 | -102.0884
18.4191 | -102.0819
18.4229 | -102.0733
18.4266 | -102.0652
18.4304 | -102.059
18.4341 | -102.0548
18.4379 | -102.0511
18.4416 | -102.0479
18.4454 | -102.0457
18.4491 | -102.0463
18.4529 | -102.0455
18.4566 | -102.0403
18.4604 | -102.0368
18.4641 | -102.0359
18.4679 | -102.0371
18.4716 | -102.0391
18.4754 | -102.0403
18.4791 | -102.0409
18.4829 | -102.0406
18.4866 | -102.0436
18.4904 | -102.0457
18.4941 | -102.0454
18.4979 | -102.0453
18.5016 | -102.0455
18.5054 | -102.0451
18.5091 | -102.0473
18.5129 | -102.052
18.5166 | -102.0566
18.5204 | -102.0614
18.5241 | -102.068
18.5279 | -102.074
18.5316 | -102.0799
18.5354 | -102.0873
18.5391 | -102.0941
18.5429 | -102.0993
18.5466 | -102.1052
18.5504 | -102.1124
18.5541 | -102.1212
18.5579 | -102.1297
18.5617 | -102.1355
18.5654 | -102.1421
18.5692 | -102.1486
18.5729 | -102.1521
18.5767 | -102.1557
18.5804 | -102.1599
18.5842 | -102.1624
18.5879 | -102.162
18.5917 | -102.1619
18.5954 | -102.1606
18.5992 | -102.1595
18.6029 | -102.1572
18.6067 | -102.1535
18.6104 | -102.1487
18.6142 | -102.1447
18.6179 | -102.1413
18.6217 | -102.1383
18.6254 | -102.133
18.6292 | -102.1292
18.6329 | -102.1285
18.6367 | -102.1283
18.6404 | -102.128
18.6442 | -102.124
18.6479 | -102.1196
18.6517 | -102.1168
18.6554 | -102.114
18.6592 | -102.112
18.6629 | -102.1107
18.6667 | -102.1127
18.6704 | -102.1143
18.6742 | -102.1152
18.6779 | -102.1171
18.6817 | -102.1187
18.6854 | -102.1191
18.6892 | -102.1209
18.6929 | -102.1254
18.6967 | -102.1314
18.7004 | -102.136
18.7042 | -102.1414
18.7079 | -102.1481
18.7117 | -102.1537
18.7154 | -102.1613
18.7192 | -102.1707
18.7229 | -102.1797
18.7267 | -102.1868
18.7304 | -102.1943
18.7342 | -102.2022
18.7379 | -102.2085
18.7417 | -102.2142
18.7454 | -102.2203
18.7492 | -102.2273
18.7529 | -102.2303
18.7567 | -102.2345
18.7604 | -102.2408
18.7642 | -102.2469
18.7679 | -102.2514
18.7717 | -102.2519
18.7754 | -102.2499
18.7792 | -102.2503
18.7829 | -102.2482
18.7867 | -102.2426
18.7904 | -102.2354
18.7942 | -102.2319
18.7979 | -102.2272
18.8017 | -102.2179
18.8054 | -102.209
18.8092 | -102.2003
18.8129 | -102.1877
18.8167 | -102.1746
18.8204 | -102.1632
18.8242 | -102.1506
18.8279 | -102.138
18.8317 | -102.1236
18.8354 | -102.1099
18.8392 | -102.0987
18.8429 | -102.0853
18.8467 | -102.0699
18.8504 | -102.0545
18.8542 | -102.0391
18.8579 | -102.0228
18.8617 | -102.0043
18.8654 | -101.9868
18.8692 | -101.9703
18.8729 | -101.9552
18.8767 | -101.94
18.8804 | -101.9231
18.8842 | -101.9066
18.888 | -101.8898
18.8917 | -101.8753
18.8955 | -101.8588
18.8992 | -101.8421
18.903 | -101.8259
18.9067 | -101.81
18.9105 | -101.7968
18.9142 | -101.7831
18.918 | -101.7678
18.9217 | -101.7571
18.9255 | -101.7481
18.9292 | -101.7371
18.933 | -101.725
18.9367 | -101.7194
18.9405 | -101.709
18.9442 | -101.6995
18.948 | -101.6931
18.9517 | -101.6862
18.9555 | -101.6813
18.9592 | -101.6763
18.963 | -101.671
18.9667 | -101.666
18.9705 | -101.6638
18.9742 | -101.6642
18.978 | -101.6639
18.9817 | -101.6655
18.9855 | -101.6666
18.9892 | -101.6672
18.993 | -101.6705
18.9967 | -101.6754
19.0005 | -101.6798
19.0042 | -101.6857
19.008 | -101.6934
19.0117 | -101.702
19.0155 | -101.7091
19.0192 | -101.7146
19.023 | -101.7234
19.0267 | -101.7341
19.0305 | -101.7456
19.0342 | -101.7554
19.038 | -101.7645
19.0417 | -101.7741
19.0455 | -101.7829
19.0492 | -101.7891
19.053 | -101.797
19.0567 | -101.8066
19.0605 | -101.8155
19.0642 | -101.8204
19.068 | -101.8229
19.0717 | -101.8238
19.0755 | -101.8198
19.0792 | -101.8167
19.083 | -101.8127
19.0867 | -101.8062
19.0905 | -101.7934
19.0942 | -101.7783
19.098 | -101.7611
19.1017 | -101.7427
19.1055 | -101.721
19.1092 | -101.6993
19.113 | -101.6761
19.1167 | -101.6504
19.1205 | -101.6228
19.1242 | -101.5901
19.128 | -101.5565
19.1317 | -101.5233
19.1355 | -101.4873
19.1392 | -101.4481
19.143 | -101.4077
19.1467 | -101.368
19.1505 | -101.3301
19.1542 | -101.2931
19.158 | -101.2566
19.1617 | -101.2163
19.1655 | -101.1763
19.1692 | -101.1389
19.173 | -101.1021
19.1767 | -101.0652
19.1805 | -101.0285
19.1842 | -100.9946
19.188 | -100.9622
19.1917 | -100.9294
19.1955 | -100.898
19.1992 | -100.8692
19.203 | -100.8444
19.2068 | -100.822
19.2105 | -100.8023
19.2143 | -100.7835
19.218 | -100.7679
19.2218 | -100.7548
19.2255 | -100.7444
19.2293 | -100.7366
19.233 | -100.7323
19.2368 | -100.7304
19.2405 | -100.7329
19.2443 | -100.7385
19.248 | -100.7448
19.2518 | -100.7557
19.2555 | -100.7701
19.2593 | -100.7863
19.263 | -100.8054
19.2668 | -100.8268
19.2705 | -100.8509
19.2743 | -100.8775
19.278 | -100.9059
19.2818 | -100.9351
19.2855 | -100.968
19.2893 | -101.0053
19.293 | -101.0457
19.2968 | -101.0859
19.3005 | -101.126
19.3043 | -101.1681
19.308 | -101.2126
19.3118 | -101.2582
19.3155 | -101.3053
19.3193 | -101.3534
19.323 | -101.4017
19.3268 | -101.4508
19.3305 | -101.4989
19.3343 | -101.5464
19.338 | -101.5954
19.3418 | -101.6461
19.3455 | -101.697
19.3493 | -101.7485
19.353 | -101.7971
19.3568 | -101.8465
19.3605 | -101.895
19.3643 | -101.9453
19.368 | -101.995
19.3718 | -102.0414
19.3755 | -102.0851
19.3793 | -102.1277
19.383 | -102.1662
19.3868 | -102.2028
19.3905 | -102.2402
19.3943 | -102.2769
19.398 | -102.3144
19.4018 | -102.3479
19.4055 | -102.379
19.4093 | -102.4094
19.413 | -102.4409
19.4168 | -102.469
19.4205 | -102.4941
19.4243 | -102.5161
19.428 | -102.5364
19.4318 | -102.5554
19.4355 | -102.5783
19.4393 | -102.6017
19.443 | -102.6215
19.4468 | -102.6405
19.4505 | -102.6586
19.4543 | -102.675
19.458 | -102.693
19.4618 | -102.713
19.4655 | -102.732
19.4693 | -102.7464
19.473 | -102.763
19.4768 | -102.7758
19.4805 | -102.7864
19.4843 | -102.796
19.488 | -102.8054
19.4918 | -102.8178
19.4955 | -102.8301
19.4993 | -102.8409
19.503 | -102.8517
19.5068 | -102.8596
19.5105 | -102.8668
19.5143 | -102.8723
19.518 | -102.8764
19.5218 | -102.8778
19.5256 | -102.8809
19.5293 | -102.8833
19.5331 | -102.8829
19.5368 | -102.882
19.5406 | -102.8801
19.5443 | -102.8762
19.5481 | -102.8713
19.5518 | -102.8663
19.5556 | -102.8608
19.5593 | -102.8565
19.5631 | -102.8489
19.5668 | -102.8384
19.5706 | -102.827
19.5743 | -102.8161
19.5781 | -102.8071
19.5818 | -102.7981
19.5856 | -102.788
19.5893 | -102.7788
19.5931 | -102.7714
19.5968 | -102.7623
19.6006 | -102.7543
19.6043 | -102.7454
19.6081 | -102.7364
19.6118 | -102.7306
19.6156 | -102.725
19.6193 | -102.7188
19.6231 | -102.7121
19.6268 | -102.7075
19.6306 | -102.7032
19.6343 | -102.6995
19.6381 | -102.6956
19.6418 | -102.6927
19.6456 | -102.6921
19.6493 | -102.6938
19.6531 | -102.6966
19.6568 | -102.701
19.6606 | -102.7057
19.6643 | -102.7104
19.6681 | -102.7115
19.6718 | -102.7123
19.6756 | -102.7155
19.6793 | -102.7198
19.6831 | -102.724
19.6868 | -102.7289
19.6906 | -102.7329
19.6943 | -102.7366
19.6981 | -102.7396
19.7018 | -102.7432
19.7056 | -102.7476
19.7093 | -102.753
19.7131 | -102.7552
19.7168 | -102.7562
19.7206 | -102.7567
19.7243 | -102.7558
19.7281 | -102.7574
19.7318 | -102.76
19.7356 | -102.7595
19.7393 | -102.7568
19.7431 | -102.7563
19.7468 | -102.7564
19.7506 | -102.7557
19.7543 | -102.7549
19.7581 | -102.7534
19.7618 | -102.7541
19.7656 | -102.7543
19.7693 | -102.7538
19.7731 | -102.7513
19.7768 | -102.7488
19.7806 | -102.7473
19.7843 | -102.7449
19.7881 | -102.7416
19.7918 | -102.7385
19.7956 | -102.736
19.7993 | -102.7305
19.8031 | -102.7265
19.8068 | -102.718
19.8106 | -102.7071
19.8143 | -102.6995
19.8181 | -102.6912
19.8218 | -102.6831
19.8256 | -102.6762
19.8293 | -102.6681
19.8331 | -102.6571
19.8368 | -102.6452
19.8406 | -102.6312
19.8444 | -102.6187
19.8481 | -102.6041
19.8519 | -102.586
19.8556 | -102.5697
19.8594 | -102.5478
19.8631 | -102.5225
19.8669 | -102.4992
19.8706 | -102.477
19.8744 | -102.454
19.8781 | -102.4302
19.8819 | -102.4055
19.8856 | -102.3767
19.8894 | -102.3488
19.8931 | -102.3199
19.8969 | -102.2896
19.9006 | -102.2589
19.9044 | -102.2308
19.9081 | -102.2003
19.9119 | -102.1692
19.9156 | -102.1373
19.9194 | -102.1049
19.9231 | -102.0727
19.9269 | -102.0404
19.9306 | -102.0083
19.9344 | -101.9773
19.9381 | -101.9457
19.9419 | -101.9135
19.9456 | -101.8832
19.9494 | -101.8571
19.9531 | -101.8301
19.9569 | -101.8025
19.9606 | -101.7755
19.9644 | -101.7464
19.9681 | -101.7188
19.9719 | -101.6924
19.9756 | -101.668
19.9794 | -101.6477
19.9831 | -101.627
19.9869 | -101.6062
19.9906 | -101.5856
19.9944 | -101.5675
19.9981 | -101.549
20.0019 | -101.5332
20.0056 | -101.5216
20.0094 | -101.5086
20.0131 | -101.4985
20.0169 | -101.4982
20.0206 | -101.4867
20.0244 | -101.4761
20.0281 | -101.4679
20.0319 | -101.4628
20.0356 | -101.4586
20.0394 | -101.454
20.0431 | -101.4509
20.0469 | -101.4475
20.0506 | -101.4447
20.0544 | -101.4439
20.0581 | -101.4427
20.0619 | -101.4416
20.0656 | -101.4412
20.0694 | -101.4394
20.0731 | -101.4386
20.0769 | -101.4387
20.0806 | -101.44
20.0844 | -101.4402
20.0881 | -101.4397
20.0919 | -101.4414
20.0956 | -101.4423
20.0994 | -101.4406
20.1031 | -101.4384
20.1069 | -101.4353
20.1106 | -101.4309
20.1144 | -101.4261
20.1181 | -101.4228
20.1219 | -101.4186
20.1256 | -101.415
20.1294 | -101.4103
20.1331 | -101.408
20.1369 | -101.4068
20.1406 | -101.4023
20.1444 | -101.4019
20.1481 | -101.4015
20.1519 | -101.3978
20.1556 | -101.3948
20.1594 | -101.3941
20.1632 | -101.3941
20.1669 | -101.3947
20.1707 | -101.3986
20.1744 | -101.4031
20.1782 | -101.4056
20.1819 | -101.408
20.1857 | -101.4114
20.1894 | -101.4143
20.1932 | -101.4194
20.1969 | -101.424
20.2007 | -101.4298
20.2044 | -101.4337
20.2082 | -101.4368
20.2119 | -101.4397
20.2157 | -101.4461
20.2194 | -101.4513
20.2232 | -101.4564
20.2269 | -101.4628
20.2307 | -101.468
20.2344 | -101.4718
20.2382 | -101.4764
20.2419 | -101.4785
20.2457 | -101.4811
20.2494 | -101.4851
20.2532 | -101.4867
20.2569 | -101.4872
20.2607 | -101.4872
20.2644 | -101.4859
20.2682 | -101.4858
20.2719 | -101.4854
20.2757 | -101.4847
20.2794 | -101.484
20.2832 | -101.4839
20.2869 | -101.4827
20.2907 | -101.4782
20.2944 | -101.4725
20.2982 | -101.4654
20.3019 | -101.4567
20.3057 | -101.4514
20.3094 | -101.448
20.3132 | -101.4441
20.3169 | -101.4382
20.3207 | -101.4306
20.3244 | -101.4218
20.3282 | -101.413
20.3319 | -101.406
20.3357 | -101.3996
20.3394 | -101.3932
20.3432 | -101.3857
20.3469 | -101.3759
20.3507 | -101.3644
20.3544 | -101.3515
20.3582 | -101.3415
20.3619 | -101.3311
20.3657 | -101.32
20.3694 | -101.3081
20.3732 | -101.2972
20.3769 | -101.2864
20.3807 | -101.2763
20.3844 | -101.2646
20.3882 | -101.2501
20.3919 | -101.2369
20.3957 | -101.2214
20.3994 | -101.2069
20.4032 | -101.1935
20.4069 | -101.1799
20.4107 | -101.1662
20.4144 | -101.1499
20.4182 | -101.1328
20.4219 | -101.1182
20.4257 | -101.1026
20.4294 | -101.0846
20.4332 | -101.0657
20.4369 | -101.0483
20.4407 | -101.0343
20.4444 | -101.0173
20.4482 | -100.9976
20.4519 | -100.9733
20.4557 | -100.9507
20.4594 | -100.9303
20.4632 | -100.91
20.4669 | -100.8912
20.4707 | -100.8733
20.4744 | -100.8536
20.4782 | -100.8341
20.482 | -100.8125
20.4857 | -100.7924
20.4895 | -100.7727
20.4932 | -100.7552
20.497 | -100.7388
20.5007 | -100.7212
20.5045 | -100.7011
20.5082 | -100.6826
20.512 | -100.6652
20.5157 | -100.6481
20.5195 | -100.6313
20.5232 | -100.6148
20.527 | -100.5997
20.5307 | -100.5864
20.5345 | -100.5745
20.5382 | -100.5593
20.542 | -100.5452
20.5457 | -100.5332
20.5495 | -100.5214
20.5532 | -100.5075
20.557 | -100.495
20.5607 | -100.4879
20.5645 | -100.4799
20.5682 | -100.4732
20.572 | -100.4651
20.5757 | -100.459
20.5795 | -100.4545
20.5832 | -100.4506
20.587 | -100.448
20.5907 | -100.4427
20.5945 | -100.439
20.5982 | -100.4353
20.602 | -100.4265
20.6057 | -100.419
20.6095 | -100.4146
20.6132 | -100.4099
20.617 | -100.4027
20.6207 | -100.3958
20.6245 | -100.3878
20.6282 | -100.3779
20.632 | -100.3678
20.6357 | -100.3588
20.6395 | -100.3511
20.6432 | -100.3464
20.647 | -100.3392
20.6507 | -100.3312
20.6545 | -100.3223
20.6582 | -100.3136
20.662 | -100.3042
20.6657 | -100.2943
20.6695 | -100.2843
20.6732 | -100.275
20.677 | -100.2668
20.6807 | -100.26
20.6845 | -100.2532
20.6882 | -100.2474
20.692 | -100.2407
20.6957 | -100.2345
20.6995 | -100.2283
20.7032 | -100.2232
20.707 | -100.2215
20.7107 | -100.2214
20.7145 | -100.2207
20.7182 | -100.2226
20.722 | -100.2278
20.7257 | -100.2338
20.7295 | -100.24
20.7332 | -100.2463
20.737 | -100.2557
20.7407 | -100.2671
20.7445 | -100.2847
20.7482 | -100.3025
20.752 | -100.3195
20.7557 | -100.3354
20.7595 | -100.3535
20.7632 | -100.375
20.767 | -100.4005
20.7707 | -100.4274
20.7745 | -100.4566
20.7782 | -100.4867
20.782 | -100.5177
20.7857 | -100.5494
20.7895 | -100.5816
20.7932 | -100.615
20.797 | -100.65
20.8008 | -100.6867
20.8045 | -100.7238
20.8083 | -100.7612
20.812 | -100.799
20.8158 | -100.8367
20.8195 | -100.8727
20.8233 | -100.9087
20.827 | -100.9434
20.8308 | -100.9826
20.8345 | -101.0207
20.8383 | -101.0566
20.842 | -101.0915
20.8458 | -101.1274
20.8495 | -101.1613
20.8533 | -101.1933
20.857 | -101.2229
20.8608 | -101.2483
20.8645 | -101.2722
20.8683 | -101.2964
20.872 | -101.3182
20.8758 | -101.3372
20.8795 | -101.356
20.8833 | -101.375
20.887 | -101.3942
20.8908 | -101.4097
20.8945 | -101.4215
20.8983 | -101.4338
20.902 | -101.442
20.9058 | -101.4475
20.9095 | -101.4523
20.9133 | -101.4593
20.917 | -101.4632
20.9208 | -101.4641
20.9245 | -101.4669
20.9283 | -101.4686
20.932 | -101.4675
20.9358 | -101.4662
20.9395 | -101.4605
20.9433 | -101.4557
20.947 | -101.4491
20.9508 | -101.4442
20.9545 | -101.436
20.9583 | -101.4262
20.962 | -101.4165
20.9658 | -101.4054
20.9695 | -101.3925
20.9733 | -101.3826
20.977 | -101.3736
20.9808 | -101.3631
20.9845 | -101.3547
20.9883 | -101.3463
20.992 | -101.3363
20.9958 | -101.325
20.9995 | -101.3157
21.0033 | -101.3048
21.007 | -101.2941
21.0108 | -101.283
21.0145 | -101.2718
21.0183 | -101.26
21.022 | -101.2485
21.0258 | -101.2396
21.0295 | -101.2309
21.0333 | -101.2228
21.037 | -101.2153
21.0408 | -101.2075
21.0445 | -101.1993
21.0483 | -101.1897
21.052 | -101.1833
21.0558 | -101.1786
21.0595 | -101.1723
21.0633 | -101.1642
21.067 | -101.1559
21.0708 | -101.1488
21.0745 | -101.1425
21.0783 | -101.1381
21.082 | -101.1348
21.0858 | -101.1314
21.0895 | -101.1294
21.0933 | -101.1266
21.097 | -101.121
21.1008 | -101.121
21.1045 | -101.1171
21.1083 | -101.1142
21.112 | -101.1126
21.1158 | -101.1117
21.1196 | -101.1112
21.1233 | -101.111
21.1271 | -101.1117
21.1308 | -101.1093
21.1346 | -101.105
21.1383 | -101.1018
21.1421 | -101.0977
21.1458 | -101.0931
21.1496 | -101.0897
21.1533 | -101.0859
21.1571 | -101.0816
21.1608 | -101.0753
21.1646 | -101.0675
21.1683 | -101.0598
21.1721 | -101.0553
21.1758 | -101.0501
21.1796 | -101.044
21.1833 | -101.0382
21.1871 | -101.034
21.1908 | -101.0299
21.1946 | -101.0249
21.1983 | -101.0203
21.2021 | -101.0148
21.2058 | -101.0063
21.2096 | -100.9984
21.2133 | -100.9901
21.2171 | -100.978
21.2208 | -100.9669
21.2246 | -100.9555
21.2283 | -100.9428
21.2321 | -100.9291
21.2358 | -100.9133
21.2396 | -100.8979
21.2433 | -100.8827
21.2471 | -100.8653
21.2508 | -100.846
21.2546 | -100.8296
21.2583 | -100.8141
21.2621 | -100.7981
21.2658 | -100.7803
21.2696 | -100.7591
21.2733 | -100.7346
21.2771 | -100.7117
21.2808 | -100.6884
21.2846 | -100.665
21.2883 | -100.6427
21.2921 | -100.6208
21.2958 | -100.5982
21.2996 | -100.5745
21.3033 | -100.551
21.3071 | -100.5281
21.3108 | -100.5033
21.3146 | -100.4793
21.3183 | -100.4587
21.3221 | -100.4382
21.3258 | -100.416
21.3296 | -100.3951
21.3333 | -100.3741
21.3371 | -100.3533
21.3408 | -100.3321
21.3446 | -100.3116
21.3483 | -100.2937
21.3521 | -100.2784
21.3558 | -100.2619
21.3596 | -100.2457
21.3633 | -100.2315
21.3671 | -100.2178
21.3708 | -100.2058
21.3746 | -100.1931
21.3783 | -100.1804
21.3821 | -100.1688
21.3858 | -100.1578
21.3896 | -100.1466
21.3933 | -100.1358
21.3971 | -100.1271
21.4008 | -100.1157
21.4046 | -100.1057
21.4083 | -100.0935
21.4121 | -100.0841
21.4158 | -100.0769
21.4196 | -100.0702
21.4233 | -100.0606
21.4271 | -100.0508
21.4308 | -100.041
21.4346 | -100.0329
21.4383 | -100.0205
21.4421 | -100.0088
21.4459 | -99.9956
21.4496 | -99.9796
21.4534 | -99.9634
21.4571 | -99.9467
21.4609 | -99.9297
21.4646 | -99.9144
21.4684 | -99.9
21.4721 | -99.8819
21.4759 | -99.856
21.4796 | -99.831
21.4834 | -99.8036
21.4871 | -99.7717
21.4909 | -99.7357
21.4946 | -99.6974
21.4984 | -99.6626
21.5021 | -99.6294
21.5059 | -99.5916
21.5096 | -99.5521
21.5134 | -99.5124
21.5171 | -99.4709
21.5209 | -99.4294
21.5246 | -99.3871
21.5284 | -99.3448
21.5321 | -99.3019
21.5359 | -99.2605
21.5396 | -99.2206
21.5434 | -99.1797
21.5471 | -99.1391
21.5509 | -99.0998
21.5546 | -99.0596
21.5584 | -99.0182
21.5621 | -98.9769
21.5659 | -98.9363
21.5696 | -98.8994
21.5734 | -98.8646
21.5771 | -98.8314
21.5809 | -98.7965
21.5846 | -98.7645
21.5884 | -98.7347
21.5921 | -98.7045
21.5959 | -98.6753
21.5996 | -98.6497
21.6034 | -98.6248
21.6071 | -98.6015
21.6109 | -98.5793
21.6146 | -98.5601
21.6184 | -98.5442
21.6221 | -98.5299
21.6259 | -98.5188
21.6296 | -98.5087
21.6334 | -98.4983
21.6371 | -98.4887
21.6409 | -98.4962
21.6446 | -98.493
21.6484 | -98.4906
21.6521 | -98.4903
21.6559 | -98.4925
21.6596 | -98.4945
21.6634 | -98.4992
21.6671 | -98.5057
21.6709 | -98.514
21.6746 | -98.5234
21.6784 | -98.5348
21.6821 | -98.5487
21.6859 | -98.5629
21.6896 | -98.5773
21.6934 | -98.5911
21.6971 | -98.6062
21.7009 | -98.6232
21.7046 | -98.6405
21.7084 | -98.6587
21.7121 | -98.6764
21.7159 | -98.6957
21.7196 | -98.7127
21.7234 | -98.7299
21.7271 | -98.7463
21.7309 | -98.7638
21.7346 | -98.7816
21.7384 | -98.7989
21.7421 | -98.8153
21.7459 | -98.8341
21.7496 | -98.8499
21.7534 | -98.8671
21.7571 | -98.8856
21.7609 | -98.9021
21.7647 | -98.9168
21.7684 | -98.9307
21.7722 | -98.9429
21.7759 | -98.9541
21.7797 | -98.9665
21.7834 | -98.9803
21.7872 | -98.9903
21.7909 | -98.9997
21.7947 | -99.0103
21.7984 | -99.0205
21.8022 | -99.0288
21.8059 | -99.0325
21.8097 | -99.036
21.8134 | -99.042
21.8172 | -99.0484
21.8209 | -99.0511
21.8247 | -99.0537
21.8284 | -99.058
21.8322 | -99.0608
21.8359 | -99.0615
21.8397 | -99.0598
21.8434 | -99.0583
21.8472 | -99.0568
21.8509 | -99.0537
21.8547 | -99.0478
21.8584 | -99.0416
21.8622 | -99.0375
21.8659 | -99.0306
21.8697 | -99.0243
21.8734 | -99.0168
21.8772 | -99.0061
21.8809 | -98.9958
21.8847 | -98.9847
21.8884 | -98.9762
21.8922 | -98.9669
21.8959 | -98.9559
21.8997 | -98.9451
21.9034 | -98.9338
21.9072 | -98.9238
21.9109 | -98.9148
21.9147 | -98.9055
21.9184 | -98.8936
21.9222 | -98.8828
21.9259 | -98.8696
21.9297 | -98.8535
21.9334 | -98.8366
21.9372 | -98.8227
21.9409 | -98.8127
21.9447 | -98.8043
21.9484 | -98.7956
21.9522 | -98.7859
21.9559 | -98.7768
21.9597 | -98.7682
21.9634 | -98.7583
21.9672 | -98.7475
21.9709 | -98.7389
21.9747 | -98.7302
21.9784 | -98.7243
21.9822 | -98.7195
21.9859 | -98.7149
21.9897 | -98.7093
21.9934 | -98.7043
21.9972 | -98.6993
22.0009 | -98.696
22.0047 | -98.6949
22.0084 | -98.6933
22.0122 | -98.692
22.0159 | -98.6906
22.0197 | -98.6905
22.0234 | -98.6923
22.0272 | -98.6944
22.0309 | -98.6961
22.0347 | -98.6946
22.0384 | -98.693
22.0422 | -98.6951
22.0459 | -98.6975
22.0497 | -98.6973
22.0534 | -98.6973
22.0572 | -98.6984
22.0609 | -98.699
22.0647 | -98.7
22.0684 | -98.6989
22.0722 | -98.6948
22.0759 | -98.6935
22.0797 | -98.6926
22.0835 | -98.6924
22.0872 | -98.6923
22.091 | -98.6916
22.0947 | -98.6908
22.0985 | -98.6915
22.1022 | -98.6925
22.106 | -98.6902
22.1097 | -98.6875
22.1135 | -98.6836
22.1172 | -98.6782
22.121 | -98.6712
22.1247 | -98.6632
22.1285 | -98.6535
22.1322 | -98.6462
22.136 | -98.639
22.1397 | -98.6304
22.1435 | -98.6235
22.1472 | -98.6158
22.151 | -98.6075
22.1547 | -98.5994
22.1585 | -98.5911
22.1622 | -98.5835
22.166 | -98.5789
22.1697 | -98.5741
22.1735 | -98.5678
22.1772 | -98.5651
22.181 | -98.5618
22.1847 | -98.5649
22.1885 | -98.5631
22.1922 | -98.5636
22.196 | -98.5647
22.1997 | -98.5663
22.2035 | -98.5667
22.2072 | -98.567
22.211 | -98.5692
22.2147 | -98.5752
22.2185 | -98.5831
22.2222 | -98.5895
22.226 | -98.5959
22.2297 | -98.6047
22.2335 | -98.6167
22.2372 | -98.6277
22.241 | -98.6413
22.2447 | -98.6585
22.2485 | -98.6762
22.2522 | -98.6943
22.256 | -98.7113
22.2597 | -98.7269
22.2635 | -98.7416
22.2672 | -98.7582
22.271 | -98.776
22.2747 | -98.7954
22.2785 | -98.8171
22.2822 | -98.8383
22.286 | -98.8578
22.2897 | -98.8779
22.2935 | -98.8964
22.2972 | -98.9143
22.301 | -98.9336
22.3047 | -98.9526
22.3085 | -98.9722
22.3122 | -98.9926
22.316 | -99.0136
22.3197 | -99.0344
22.3235 | -99.0533
22.3272 | -99.0697
22.331 | -99.0858
22.3347 | -99.1009
22.3385 | -99.1156
22.3422 | -99.1297
22.346 | -99.1434
22.3497 | -99.1591
22.3535 | -99.1747
22.3572 | -99.1858
22.361 | -99.1938
22.3647 | -99.2073
22.3685 | -99.2214
22.3722 | -99.234
22.376 | -99.2465
22.3797 | -99.2566
22.3835 | -99.2664
22.3872 | -99.279
22.391 | -99.2892
22.3947 | -99.2979
22.3985 | -99.3035
22.4022 | -99.3068
22.406 | -99.3108
22.4098 | -99.3148
22.4135 | -99.3183
22.4173 | -99.3193
22.421 | -99.3224
22.4248 | -99.3255
22.4285 | -99.33
22.4323 | -99.3353
22.436 | -99.3363
22.4398 | -99.3349
22.4435 | -99.3337
22.4473 | -99.331
22.451 | -99.3263
22.4548 | -99.3224
22.4585 | -99.3181
22.4623 | -99.3127
22.466 | -99.3077
22.4698 | -99.2988
22.4735 | -99.2908
22.4773 | -99.2816
22.481 | -99.2703
22.4848 | -99.2594
22.4885 | -99.2481
22.4923 | -99.2366
22.496 | -99.2241
22.4998 | -99.2093
22.5035 | -99.1986
22.5073 | -99.1876
22.511 | -99.1734
22.5148 | -99.1592
22.5185 | -99.1455
22.5223 | -99.1308
22.526 | -99.1165
22.5298 | -99.1017
22.5335 | -99.0842
22.5373 | -99.0656
22.541 | -99.0493
22.5448 | -99.0341
22.5485 | -99.0181
22.5523 | -99.0028
22.556 | -98.9864
22.5598 | -98.9706
22.5635 | -98.9508
22.5673 | -98.9305
22.571 | -98.9105
22.5748 | -98.8934
22.5785 | -98.878
22.5823 | -98.8629
22.586 | -98.8469
22.5898 | -98.8323
22.5935 | -98.8155
22.5973 | -98.7991
22.601 | -98.7827
22.6048 | -98.7684
22.6085 | -98.7548
22.6123 | -98.7406
22.616 | -98.7283
22.6198 | -98.7185
22.6235 | -98.7085
22.6273 | -98.6983
22.631 | -98.6888
22.6348 | -98.6778
22.6385 | -98.6679
22.6423 | -98.662
22.646 | -98.6581
22.6498 | -98.6556
22.6535 | -98.654
22.6573 | -98.6517
22.661 | -98.6489
22.6648 | -98.6491
22.6685 | -98.6494
22.6723 | -98.6519
22.676 | -98.6556
22.6798 | -98.6619
22.6835 | -98.6663
22.6873 | -98.6717
22.691 | -98.6791
22.6948 | -98.6868
22.6985 | -98.697
22.7023 | -98.7079
22.706 | -98.7187
22.7098 | -98.7312
22.7135 | -98.7456
22.7173 | -98.7604
22.7211 | -98.7766
22.7248 | -98.7962
22.7286 | -98.8122
22.7323 | -98.8289
22.7361 | -98.8466
22.7398 | -98.866
22.7436 | -98.8818
22.7473 | -98.8985
22.7511 | -98.9141
22.7548 | -98.9287
22.7586 | -98.9446
22.7623 | -98.9618
22.7661 | -98.978
22.7698 | -98.9931
22.7736 | -99.0087
22.7773 | -99.0231
22.7811 | -99.0331
22.7848 | -99.0439
22.7886 | -99.055
22.7923 | -99.0667
22.7961 | -99.0737
22.7998 | -99.0804
22.8036 | -99.0871
22.8073 | -99.0945
22.8111 | -99.0991
22.8148 | -99.1009
22.8186 | -99.1034
22.8223 | -99.1039
22.8261 | -99.1003
22.8298 | -99.0953
22.8336 | -99.0899
22.8373 | -99.0829
22.8411 | -99.0759
22.8448 | -99.0685
22.8486 | -99.06
22.8523 | -99.052
22.8561 | -99.0432
22.8598 | -99.0337
22.8636 | -99.0218
22.8673 | -99.009
22.8711 | -98.9928
22.8748 | -98.9765
22.8786 | -98.9618
22.8823 | -98.9452
22.8861 | -98.9316
22.8898 | -98.9182
22.8936 | -98.9025
22.8973 | -98.8826
22.9011 | -98.8636
22.9048 | -98.8446
22.9086 | -98.8252
22.9123 | -98.8071
22.9161 | -98.7896
22.9198 | -98.773
22.9236 | -98.7568
22.9273 | -98.7393
22.9311 | -98.7228
22.9348 | -98.7069
22.9386 | -98.6897
22.9423 | -98.6719
22.9461 | -98.6555
22.9498 | -98.6393
22.9536 | -98.622
22.9573 | -98.6071
22.9611 | -98.5929
22.9648 | -98.5781
22.9686 | -98.5637
22.9723 | -98.5472
22.9761 | -98.5294
22.9798 | -98.5142
22.9836 | -98.4997
22.9873 | -98.4846
22.9911 | -98.4694
22.9948 | -98.4544
22.9986 | -98.4388
23.0023 | -98.4215
23.0061 | -98.4056
23.0098 | -98.392
23.0136 | -98.3776
23.0173 | -98.3653
23.0211 | -98.3519
23.0248 | -98.3411
23.0286 | -98.3293
23.0323 | -98.3175
23.0361 | -98.3057
23.0399 | -98.2935
23.0436 | -98.2833
23.0474 | -98.2741
23.0511 | -98.2634
23.0549 | -98.2518
23.0586 | -98.2405
23.0624 | -98.2317
23.0661 | -98.224
23.0699 | -98.2172
23.0736 | -98.2072
23.0774 | -98.1983
23.0811 | -98.1896
23.0849 | -98.1818
23.0886 | -98.176
23.0924 | -98.1705
23.0961 | -98.1669
23.0999 | -98.1641
23.1036 | -98.1622
23.1074 | -98.1617
23.1111 | -98.1616
23.1149 | -98.1629
23.1186 | -98.1647
23.1224 | -98.1688
23.1261 | -98.1725
23.1299 | -98.1754
23.1336 | -98.1786
23.1374 | -98.1838
23.1411 | -98.1901
23.1449 | -98.1962
23.1486 | -98.2057
23.1524 | -98.2163
23.1561 | -98.2275
23.1599 | -98.2405
23.1636 | -98.2528
23.1674 | -98.2666
23.1711 | -98.2812
23.1749 | -98.2973
23.1786 | -98.3121
23.1824 | -98.3281
23.1861 | -98.3462
23.1899 | -98.366
23.1936 | -98.3852
23.1974 | -98.4066
23.2011 | -98.4286
23.2049 | -98.4509
23.2086 | -98.4717
23.2124 | -98.4892
23.2161 | -98.5051
23.2199 | -98.5236
23.2236 | -98.5441
23.2274 | -98.561
23.2311 | -98.5764
23.2349 | -98.5905
23.2386 | -98.6027
23.2424 | -98.6151
23.2461 | -98.6255
23.2499 | -98.6365
23.2536 | -98.6464
23.2574 | -98.6538
23.2611 | -98.6592
23.2649 | -98.6518
23.2686 | -98.6558
23.2724 | -98.6585
23.2761 | -98.6584
23.2799 | -98.6593
23.2836 | -98.6606
23.2874 | -98.6567
23.2911 | -98.6483
23.2949 | -98.64
23.2986 | -98.6314
23.3024 | -98.6238
23.3061 | -98.6137
23.3099 | -98.5992
23.3136 | -98.584
23.3174 | -98.5661
23.3211 | -98.5495
23.3249 | -98.5334
23.3286 | -98.5166
23.3324 | -98.4965
23.3361 | -98.4743
23.3399 | -98.4496
23.3436 | -98.4247
23.3474 | -98.4005
23.3511 | -98.3765
23.3549 | -98.3523
23.3587 | -98.3314
23.3624 | -98.3092
23.3662 | -98.2888
23.3699 | -98.2702
23.3737 | -98.249
23.3774 | -98.2291
23.3812 | -98.2111
23.3849 | -98.1954
23.3887 | -98.1797
23.3924 | -98.1624
23.3962 | -98.147
23.3999 | -98.1346
23.4037 | -98.1247
23.4074 | -98.1148
23.4112 | -98.1059
23.4149 | -98.0977
23.4187 | -98.0889
23.4224 | -98.0813
23.4262 | -98.0747
23.4299 | -98.0705
23.4337 | -98.0667
23.4374 | -98.0649
23.4412 | -98.0658
23.4449 | -98.0659
23.4487 | -98.0667
23.4524 | -98.0699
23.4562 | -98.0763
23.4599 | -98.0843
23.4637 | -98.092
23.4674 | -98.101
23.4712 | -98.1106
23.4749 | -98.1216
23.4787 | -98.1335
23.4824 | -98.1478
23.4862 | -98.1639
23.4899 | -98.1806
23.4937 | -98.1972
23.4974 | -98.213
23.5012 | -98.23
23.5049 | -98.2423
23.5087 | -98.2567
23.5124 | -98.2737
23.5162 | -98.2879
23.5199 | -98.301
23.5237 | -98.3173
23.5274 | -98.3354
23.5312 | -98.3515
23.5349 | -98.3657
23.5387 | -98.3816
23.5424 | -98.4004
23.5462 | -98.4178
23.5499 | -98.4311
23.5537 | -98.4417
23.5574 | -98.454
23.5612 | -98.4673
23.5649 | -98.4783
23.5687 | -98.4881
23.5724 | -98.4976
23.5762 | -98.505
23.5799 | -98.51
23.5837 | -98.514
23.5874 | -98.516
23.5912 | -98.5167
23.5949 | -98.5143
23.5987 | -98.5113
23.6024 | -98.506
23.6062 | -98.4989
23.6099 | -98.493
23.6137 | -98.486
23.6174 | -98.4786
23.6212 | -98.473
23.6249 | -98.466
23.6287 | -98.455
23.6324 | -98.4432
23.6362 | -98.4305
23.6399 | -98.4176
23.6437 | -98.4051
23.6474 | -98.3946
23.6512 | -98.385
23.6549 | -98.3729
23.6587 | -98.3594
23.6624 | -98.348
23.6662 | -98.3395
23.6699 | -98.3304
23.6737 | -98.3204
23.6775 | -98.3106
23.6812 | -98.2999
23.685 | -98.2894
23.6887 | -98.2802
23.6925 | -98.2697
23.6962 | -98.2618
23.7 | -98.2533
23.7037 | -98.247
23.7075 | -98.2425
23.7112 | -98.2375
23.715 | -98.2343
23.7187 | -98.233
23.7225 | -98.2317
23.7262 | -98.2321
23.73 | -98.2341
23.7337 | -98.2367
23.7375 | -98.2399
23.7412 | -98.2448
23.745 | -98.2506
23.7487 | -98.2572
23.7525 | -98.2646
23.7562 | -98.2742
23.76 | -98.2833
23.7637 | -98.2922
23.7675 | -98.303
23.7712 | -98.3161
23.775 | -98.3292
23.7787 | -98.3439
23.7825 | -98.3591
23.7862 | -98.3755
23.79 | -98.3908
23.7937 | -98.4051
23.7975 | -98.4223
23.8012 | -98.4418
23.805 | -98.467
23.8087 | -98.4868
23.8125 | -98.5064
23.8162 | -98.5259
23.82 | -98.5459
23.8237 | -98.5661
23.8275 | -98.5865
23.8312 | -98.61
23.835 | -98.633
23.8387 | -98.6544
23.8425 | -98.6802
23.8462 | -98.7054
23.85 | -98.728
23.8537 | -98.7514
23.8575 | -98.7754
23.8612 | -98.7955
23.865 | -98.8119
23.8687 | -98.8301
23.8725 | -98.8497
23.8762 | -98.8681
23.88 | -98.8857
23.8837 | -98.9023
23.8875 | -98.9186
23.8912 | -98.9359
23.895 | -98.951
23.8987 | -98.9645
23.9025 | -98.9772
23.9062 | -98.9882
23.91 | -98.9983
23.9137 | -99.0065
23.9175 | -99.0145
23.9212 | -99.0217
23.925 | -99.0297
23.9287 | -99.0393
23.9325 | -99.0495
23.9362 | -99.0582
23.94 | -99.066
23.9437 | -99.0726
23.9475 | -99.0787
23.9512 | -99.0822
23.955 | -99.086
23.9587 | -99.0908
23.9625 | -99.0982
23.9662 | -99.1051
23.97 | -99.1101
23.9737 | -99.1154
23.9775 | -99.1229
23.9812 | -99.1313
23.985 | -99.1381
23.9887 | -99.1439
23.9925 | -99.1499
23.9962 | -99.156
24 | -99.1627
24.0038 | -99.1709
24.0075 | -99.1789
24.0113 | -99.1883
24.015 | -99.1985
24.0188 | -99.2094
24.0225 | -99.2203
24.0263 | -99.2298
24.03 | -99.2411
24.0338 | -99.2524
24.0375 | -99.2644
24.0413 | -99.2748
24.045 | -99.2883
24.0488 | -99.3025
24.0525 | -99.316
24.0563 | -99.3297
24.06 | -99.344
24.0638 | -99.3596
24.0675 | -99.3733
24.0713 | -99.3865
24.075 | -99.3999
24.0788 | -99.4137
24.0825 | -99.4272
24.0863 | -99.4414
24.09 | -99.4574
24.0938 | -99.4707
24.0975 | -99.4853
24.1013 | -99.4988
24.105 | -99.5111
24.1088 | -99.5232
24.1125 | -99.5353
24.1163 | -99.5483
24.12 | -99.5613
24.1238 | -99.5723
24.1275 | -99.5843
24.1313 | -99.5954
24.135 | -99.6043
24.1388 | -99.6109
24.1425 | -99.6175
24.1463 | -99.6256
24.15 | -99.6327
24.1538 | -99.64
24.1575 | -99.6478
24.1613 | -99.6534
24.165 | -99.6597
24.1688 | -99.6653
24.1725 | -99.6699
24.1763 | -99.675
24.18 | -99.6826
24.1838 | -99.6888
24.1875 | -99.6922
24.1913 | -99.6951
24.195 | -99.6989
24.1988 | -99.702
24.2025 | -99.7077
24.2063 | -99.7154
24.21 | -99.7207
24.2138 | -99.7265
24.2175 | -99.7311
24.2213 | -99.7351
24.225 | -99.7423
24.2288 | -99.7476
24.2325 | -99.7509
24.2363 | -99.7569
24.24 | -99.7629
24.2438 | -99.7692
24.2475 | -99.7744
24.2513 | -99.78
24.255 | -99.7879
24.2588 | -99.7944
24.2625 | -99.7996
24.2663 | -99.7938
24.27 | -99.7812
24.2738 | -99.767
24.2775 | -99.7517
24.2813 | -99.7319
24.285 | -99.7162
24.2888 | -99.7037
24.2925 | -99.6915
24.2963 | -99.6775
24.3 | -99.6601
24.3038 | -99.6418
24.3075 | -99.6231
24.3113 | -99.6016
24.315 | -99.5773
24.3188 | -99.5529
24.3225 | -99.5267
24.3263 | -99.4967
24.3301 | -99.4662
24.3338 | -99.4353
24.3376 | -99.4023
24.3413 | -99.3669
24.3451 | -99.3299
24.3488 | -99.2823
24.3526 | -99.2467
24.3563 | -99.2095
24.3601 | -99.1721
24.3638 | -99.1324
24.3676 | -99.0948
24.3713 | -99.0583
24.3751 | -99.0209
24.3788 | -98.984
24.3826 | -98.9475
24.3863 | -98.9131
24.3901 | -98.8781
24.3938 | -98.8448
24.3976 | -98.8127
24.4013 | -98.7821
24.4051 | -98.7552
24.4088 | -98.7291
24.4126 | -98.704
24.4163 | -98.6841
24.4201 | -98.668
24.4238 | -98.653
24.4276 | -98.6408
24.4313 | -98.63
24.4351 | -98.6207
24.4388 | -98.6163
24.4426 | -98.6152
24.4463 | -98.6174
24.4501 | -98.623
24.4538 | -98.6308
24.4576 | -98.6404
24.4613 | -98.6534
24.4651 | -98.6696
24.4688 | -98.6887
24.4726 | -98.7114
24.4763 | -98.7382
24.4801 | -98.7668
24.4838 | -98.7978
24.4876 | -98.8302
24.4913 | -98.865
24.4951 | -98.9027
24.4988 | -98.9424
24.5026 | -98.9829
24.5063 | -99.0253
24.5101 | -99.0686
24.5138 | -99.1143
24.5176 | -99.161
24.5213 | -99.2076
24.5251 | -99.2536
24.5288 | -99.3011
24.5326 | -99.3503
24.5363 | -99.4004
24.5401 | -99.4497
24.5438 | -99.4975
24.5476 | -99.5443
24.5513 | -99.5908
24.5551 | -99.6364
24.5588 | -99.6765
24.5626 | -99.7125
24.5663 | -99.7475
24.5701 | -99.7764
24.5738 | -99.7996
24.5776 | -99.826
24.5813 | -99.8492
24.5851 | -99.8685
24.5888 | -99.8809
24.5926 | -99.8859
24.5963 | -99.8925
24.6001 | -99.9005
24.6038 | -99.9072
24.6076 | -99.9136
24.6113 | -99.926
24.6151 | -99.9373
24.6188 | -99.9422
24.6226 | -99.9412
24.6263 | -99.9403
24.6301 | -99.9285
24.6338 | -99.9079
24.6376 | -99.8832
24.6413 | -99.8561
24.6451 | -99.8278
24.6489 | -99.797
24.6526 | -99.7658
24.6564 | -99.7352
24.6601 | -99.7019
24.6639 | -99.6638
24.6676 | -99.6242
24.6714 | -99.5865
24.6751 | -99.549
24.6789 | -99.5124
24.6826 | -99.4761
24.6864 | -99.4391
24.6901 | -99.4025
24.6939 | -99.3645
24.6976 | -99.327
24.7014 | -99.2924
24.7051 | -99.26
24.7089 | -99.2276
24.7126 | -99.1996
24.7164 | -99.1728
24.7201 | -99.1472
24.7239 | -99.124
24.7276 | -99.1021
24.7314 | -99.0837
24.7351 | -99.0679
24.7389 | -99.0556
24.7426 | -99.0464
24.7464 | -99.0394
24.7501 | -99.0365
24.7539 | -99.0357
24.7576 | -99.0384
24.7614 | -99.0461
24.7651 | -99.0555
24.7689 | -99.0669
24.7726 | -99.0812
24.7764 | -99.0998
24.7801 | -99.1224
24.7839 | -99.1482
24.7876 | -99.1763
24.7914 | -99.2077
24.7951 | -99.2428
24.7989 | -99.2799
24.8026 | -99.3192
24.8064 | -99.3614
24.8101 | -99.4052
24.8139 | -99.4482
24.8176 | -99.4924
24.8214 | -99.5394
24.8251 | -99.5881
24.8289 | -99.6389
24.8326 | -99.6915
24.8364 | -99.7454
24.8401 | -99.7999
24.8439 | -99.8545
24.8476 | -99.9089
24.8514 | -99.9635
24.8551 | -100.0171
24.8589 | -100.069
24.8626 | -100.1202
24.8664 | -100.1723
24.8701 | -100.2239
24.8739 | -100.2738
24.8776 | -100.3205
24.8814 | -100.3653
24.8851 | -100.4063
24.8889 | -100.4318
24.8926 | -100.4717
24.8964 | -100.5102
24.9001 | -100.5478
24.9039 | -100.5823
24.9076 | -100.6154
24.9114 | -100.6505
24.9151 | -100.6874
24.9189 | -100.722
24.9226 | -100.7554
24.9264 | -100.7923
24.9301 | -100.8193
24.9339 | -100.8417
24.9376 | -100.8607
24.9414 | -100.8815
24.9451 | -100.9067
24.9489 | -100.9214
24.9526 | -100.9276
24.9564 | -100.9327
24.9601 | -100.9346
24.9639 | -100.9337
24.9677 | -100.9343
24.9714 | -100.9352
24.9752 | -100.9368
24.9789 | -100.9377
24.9827 | -100.9393
24.9864 | -100.9393
24.9902 | -100.9387
24.9939 | -100.9373
24.9977 | -100.9359
25.0014 | -100.9351
25.0052 | -100.9337
25.0089 | -100.9311
25.0127 | -100.9304
25.0164 | -100.9319
25.0202 | -100.9319
25.0239 | -100.9303
25.0277 | -100.9257
25.0314 | -100.9184
25.0352 | -100.9126
25.0389 | -100.9113
25.0427 | -100.9103
25.0464 | -100.9103
25.0502 | -100.9097
25.0539 | -100.9071
25.0577 | -100.9042
25.0614 | -100.9022
25.0652 | -100.8978
25.0689 | -100.8915
25.0727 | -100.8854
25.0764 | -100.882
25.0802 | -100.8798
25.0839 | -100.8772
25.0877 | -100.8728
25.0914 | -100.8682
25.0952 | -100.8642
25.0989 | -100.8616
25.1027 | -100.859
25.1064 | -100.8567
25.1102 | -100.8563
25.1139 | -100.8546
25.1177 | -100.8502
25.1214 | -100.8473
25.1252 | -100.8467
25.1289 | -100.846
25.1327 | -100.8429
25.1364 | -100.8387
25.1402 | -100.835
25.1439 | -100.8309
25.1477 | -100.8275
25.1514 | -100.8252
25.1552 | -100.8231
25.1589 | -100.8211
25.1627 | -100.8201
25.1664 | -100.8212
25.1702 | -100.8215
25.1739 | -100.8201
25.1777 | -100.8178
25.1814 | -100.8124
25.1852 | -100.8044
25.1889 | -100.7983
25.1927 | -100.7923
25.1964 | -100.787
25.2002 | -100.7833
25.2039 | -100.7768
25.2077 | -100.7713
25.2114 | -100.7681
25.2152 | -100.7647
25.2189 | -100.7593
25.2227 | -100.7529
25.2264 | -100.7494
25.2302 | -100.7464
25.2339 | -100.7407
25.2377 | -100.7359
25.2414 | -100.7267
25.2452 | -100.7195
25.2489 | -100.7119
25.2527 | -100.703
25.2564 | -100.6928
25.2602 | -100.6816
25.2639 | -100.6663
25.2677 | -100.6524
25.2714 | -100.641
25.2752 | -100.6291
25.2789 | -100.6172
25.2827 | -100.6031
25.2865 | -100.5894
25.2902 | -100.5763
25.294 | -100.5624
25.2977 | -100.5476
25.3015 | -100.5333
25.3052 | -100.5183
25.309 | -100.5016
25.3127 | -100.4856
25.3165 | -100.4692
25.3202 | -100.4543
25.324 | -100.4381
25.3277 | -100.4208
25.3315 | -100.4034
25.3352 | -100.3865
25.339 | -100.3672
25.3427 | -100.3492
25.3465 | -100.3321
25.3502 | -100.3161
25.354 | -100.3007
25.3577 | -100.2836
25.3615 | -100.2667
25.3652 | -100.2478
25.369 | -100.2273
25.3727 | -100.2067
25.3765 | -100.1866
25.3802 | -100.1685
25.384 | -100.1514
25.3877 | -100.1355
25.3915 | -100.1193
25.3952 | -100.1022
25.399 | -100.0849
25.4027 | -100.0678
25.4065 | -100.053
25.4102 | -100.0383
25.414 | -100.0233
25.4177 | -100.0082
25.4215 | -99.992
25.4252 | -99.9773
25.429 | -99.961
25.4327 | -99.9464
25.4365 | -99.9298
25.4402 | -99.9118
25.444 | -99.8967
25.4477 | -99.8804
25.4515 | -99.8615
25.4552 | -99.839
25.459 | -99.819
25.4627 | -99.8014
25.4665 | -99.7794
25.4702 | -99.7576
25.474 | -99.7386
25.4777 | -99.7197
25.4815 | -99.698
25.4852 | -99.6748
25.489 | -99.6518
25.4927 | -99.6284
25.4965 | -99.6029
25.5002 | -99.5757
25.504 | -99.5497
25.5077 | -99.5237
25.5115 | -99.4975
25.5152 | -99.4706
25.519 | -99.4426
25.5227 | -99.4137
25.5265 | -99.3844
25.5302 | -99.3549
25.534 | -99.3259
25.5377 | -99.298
25.5415 | -99.2705
25.5452 | -99.2435
25.549 | -99.2154
25.5527 | -99.1867
25.5565 | -99.1576
25.5602 | -99.128
25.564 | -99.1012
25.5677 | -99.0751
25.5715 | -99.0483
25.5752 | -99.023
25.579 | -98.9988
25.5827 | -98.9765
25.5865 | -98.9535
25.5902 | -98.9317
25.594 | -98.9124
25.5978 | -98.8944
25.6015 | -98.8783
25.6053 | -98.8637
25.609 | -98.8506
25.6128 | -98.8387
25.6165 | -98.8272
25.6203 | -98.819
25.624 | -98.8157
25.6278 | -98.811
25.6315 | -98.8059
25.6353 | -98.8019
25.639 | -98.8002
25.6428 | -98.8016
25.6465 | -98.8035
25.6503 | -98.8069
25.654 | -98.8135
25.6578 | -98.823
25.6615 | -98.8323
25.6653 | -98.8417
25.669 | -98.8532
25.6728 | -98.8669
25.6765 | -98.8824
25.6803 | -98.8989
25.684 | -98.9168
25.6878 | -98.9363
25.6915 | -98.9552
25.6953 | -98.9742
25.699 | -98.9963
25.7028 | -99.0196
25.7065 | -99.0438
25.7103 | -99.0678
25.714 | -99.0933
25.7178 | -99.1195
25.7215 | -99.1457
25.7253 | -99.174
25.729 | -99.2042
25.7328 | -99.2341
25.7365 | -99.2641
25.7403 | -99.2916
25.744 | -99.3196
25.7478 | -99.3471
25.7515 | -99.3736
25.7553 | -99.3997
25.759 | -99.4257
25.7628 | -99.4522
25.7665 | -99.4758
25.7703 | -99.4999
25.774 | -99.5247
25.7778 | -99.5481
25.7815 | -99.5732
25.7853 | -99.5976
25.789 | -99.6217
25.7928 | -99.6458
25.7965 | -99.6677
25.8003 | -99.6891
25.804 | -99.7119
25.8078 | -99.7319
25.8115 | -99.7532
25.8153 | -99.7748
25.819 | -99.7976
25.8228 | -99.8162
25.8265 | -99.8326
25.8303 | -99.852
25.834 | -99.8695
25.8378 | -99.8852
25.8415 | -99.9026
25.8453 | -99.9194
25.849 | -99.9349
25.8528 | -99.9496
25.8565 | -99.9663
25.8603 | -99.984
25.864 | -100.0001
25.8678 | -100.0122
25.8715 | -100.0244
25.8753 | -100.0386
25.879 | -100.0525
25.8828 | -100.0665
25.8865 | -100.0812
25.8903 | -100.0971
25.894 | -100.1138
25.8978 | -100.1272
25.9015 | -100.1388
25.9053 | -100.1532
25.909 | -100.1692
25.9128 | -100.1828
25.9165 | -100.1967
25.9203 | -100.211
25.9241 | -100.2245
25.9278 | -100.2372
25.9316 | -100.2472
25.9353 | -100.2555
25.9391 | -100.2643
25.9428 | -100.2721
25.9466 | -100.2778
25.9503 | -100.2838
25.9541 | -100.2902
25.9578 | -100.2947
25.9616 | -100.2972
25.9653 | -100.2999
25.9691 | -100.289
25.9728 | -100.2858
25.9766 | -100.2809
25.9803 | -100.2768
25.9841 | -100.2743
25.9878 | -100.2678
25.9916 | -100.2576
25.9953 | -100.2485
25.9991 | -100.2376
26.0028 | -100.2265
26.0066 | -100.2137
26.0103 | -100.1996
26.0141 | -100.1847
26.0178 | -100.1675
26.0216 | -100.1501
26.0253 | -100.1307
26.0291 | -100.109
26.0328 | -100.0889
26.0366 | -100.068
26.0403 | -100.0437
26.0441 | -100.0178
26.0478 | -99.9909
26.0516 | -99.9639
26.0553 | -99.9375
26.0591 | -99.9139
26.0628 | -99.8885
26.0666 | -99.8593
26.0703 | -99.8292
26.0741 | -99.7988
26.0778 | -99.7703
26.0816 | -99.7412
26.0853 | -99.7137
26.0891 | -99.6874
26.0928 | -99.6616
26.0966 | -99.6353
26.1003 | -99.6088
26.1041 | -99.585
26.1078 | -99.5618
26.1116 | -99.5386
26.1153 | -99.5172
26.1191 | -99.4971
26.1228 | -99.4765
26.1266 | -99.4565
26.1303 | -99.4388
26.1341 | -99.4221
26.1378 | -99.4055
26.1416 | -99.3906
26.1453 | -99.3771
26.1491 | -99.3623
26.1528 | -99.3496
26.1566 | -99.3389
26.1603 | -99.3289
26.1641 | -99.3222
26.1678 | -99.3158
26.1716 | -99.3097
26.1753 | -99.3046
26.1791 | -99.2994
26.1828 | -99.296
26.1866 | -99.2939
26.1903 | -99.292
26.1941 | -99.2909
26.1978 | -99.2917
26.2016 | -99.2915
26.2053 | -99.2928
26.2091 | -99.2935
26.2128 | -99.2966
26.2166 | -99.3005
26.2203 | -99.303
26.2241 | -99.3065
26.2278 | -99.3127
26.2316 | -99.3196
26.2353 | -99.3268
26.2391 | -99.3336
26.2429 | -99.3393
26.2466 | -99.3449
26.2504 | -99.3519
26.2541 | -99.3596
26.2579 | -99.3667
26.2616 | -99.3725
26.2654 | -99.3805
26.2691 | -99.3897
26.2729 | -99.3975
26.2766 | -99.4056
26.2804 | -99.4151
26.2841 | -99.4236
26.2879 | -99.4305
26.2916 | -99.439
26.2954 | -99.4461
26.2991 | -99.4541
26.3029 | -99.4647
26.3066 | -99.4735
26.3104 | -99.4797
26.3141 | -99.4852
26.3179 | -99.4897
26.3216 | -99.4956
26.3254 | -99.5033
26.3291 | -99.5116
26.3329 | -99.5175
26.3366 | -99.5229
26.3404 | -99.5301
26.3441 | -99.536
26.3479 | -99.5404
26.3516 | -99.5442
26.3554 | -99.5448
26.3591 | -99.5444
26.3629 | -99.5432
26.3666 | -99.5436
26.3704 | -99.5446
26.3741 | -99.5414
26.3779 | -99.5388
26.3816 | -99.5359
26.3854 | -99.5324
26.3891 | -99.5286
26.3929 | -99.5235
26.3966 | -99.5181
26.4004 | -99.5118
26.4041 | -99.5037
26.4079 | -99.4895
26.4116 | -99.477
26.4154 | -99.4675
26.4191 | -99.4554
26.4229 | -99.4418
26.4266 | -99.4269
26.4304 | -99.4107
26.4341 | -99.3958
26.4379 | -99.3827
26.4416 | -99.368
26.4454 | -99.3526
26.4491 | -99.3354
26.4529 | -99.3183
26.4566 | -99.3014
26.4604 | -99.2827
26.4641 | -99.2627
26.4679 | -99.2443
26.4716 | -99.2262
26.4754 | -99.2066
26.4791 | -99.188
26.4829 | -99.1716
26.4866 | -99.1539
26.4904 | -99.1359
26.4941 | -99.1161
26.4979 | -99.0953
26.5016 | -99.0766
26.5054 | -99.0614
26.5091 | -99.0445
26.5129 | -99.0318
26.5166 | -99.0169
26.5204 | -99.0039
26.5241 | -98.993
26.5279 | -98.9809
26.5316 | -98.9686
26.5354 | -98.9573
26.5391 | -98.946
26.5429 | -98.935
26.5466 | -98.925
26.5504 | -98.9149
26.5541 | -98.9061
26.5579 | -98.8975
26.5617 | -98.8897
26.5654 | -98.8836
26.5692 | -98.878
26.5729 | -98.8714
26.5767 | -98.8674
26.5804 | -98.8625
26.5842 | -98.8596
26.5879 | -98.8552
26.5917 | -98.852
26.5954 | -98.8494
26.5992 | -98.8465
26.6029 | -98.8421
26.6067 | -98.8389
26.6104 | -98.8375
26.6142 | -98.8363
26.6179 | -98.8338
26.6217 | -98.8315
26.6254 | -98.8297
26.6292 | -98.8252
26.6329 | -98.8202
26.6367 | -98.8164
26.6404 | -98.8124
26.6442 | -98.8061
26.6479 | -98.8009
26.6517 | -98.7954
26.6554 | -98.7903
26.6592 | -98.7856
26.6629 | -98.7828
26.6667 | -98.78
26.6704 | -98.7748
26.6742 | -98.7699
26.6779 | -98.7653
26.6817 | -98.7589
26.6854 | -98.7524
26.6892 | -98.7441
26.6929 | -98.7356
26.6967 | -98.7276
26.7004 | -98.7192
26.7042 | -98.7096
26.7079 | -98.7018
26.7117 | -98.6946
26.7154 | -98.686
26.7192 | -98.678
26.7229 | -98.6705
26.7267 | -98.6618
26.7304 | -98.6538
26.7342 | -98.6457
26.7379 | -98.6363
26.7417 | -98.6238
26.7454 | -98.61
26.7492 | -98.5973
26.7529 | -98.5846
26.7567 | -98.5712
26.7604 | -98.5562
26.7642 | -98.5421
26.7679 | -98.528
26.7717 | -98.5134
26.7754 | -98.498
26.7792 | -98.4806
26.7829 | -98.4671
26.7867 | -98.4541
26.7904 | -98.4347
26.7942 | -98.4168
26.7979 | -98.4
26.8017 | -98.3774
26.8054 | -98.3539
26.8092 | -98.3331
26.8129 | -98.3109
26.8167 | -98.2905
26.8204 | -98.2698
26.8242 | -98.2515
26.8279 | -98.2274
26.8317 | -98.2057
26.8354 | -98.1862
26.8392 | -98.166
26.8429 | -98.1448
26.8467 | -98.1261
26.8504 | -98.1063
26.8542 | -98.0853
26.8579 | -98.0629
26.8617 | -98.0422
26.8654 | -98.0244
26.8692 | -98.0085
26.8729 | -97.9926
26.8767 | -97.9746
26.8804 | -97.9543
26.8842 | -97.933
26.888 | -97.9119
26.8917 | -97.8906
26.8955 | -97.8689
26.8992 | -97.8486
26.903 | -97.8286
26.9067 | -97.81
26.9105 | -97.7918
26.9142 | -97.7724
26.918 | -97.7523
26.9217 | -97.7346
26.9255 | -97.7181
26.9292 | -97.7032
26.933 | -97.6887
26.9367 | -97.6754
26.9405 | -97.6618
26.9442 | -97.6474
26.948 | -97.6345
26.9517 | -97.623
26.9555 | -97.6124
26.9592 | -97.603
26.963 | -97.5935
26.9667 | -97.5885
26.9705 | -97.5841
26.9742 | -97.5807
26.978 | -97.5784
26.9817 | -97.5779
26.9855 | -97.5787
26.9892 | -97.5803
26.993 | -97.5836
26.9967 | -97.5887
27.0005 | -97.5953
27.0042 | -97.6034
27.008 | -97.6156
27.0117 | -97.63
27.0155 | -97.6451
27.0192 | -97.6614
27.023 | -97.6799
27.0267 | -97.7
27.0305 | -97.7218
27.0342 | -97.7451
27.038 | -97.7697
27.0417 | -97.7948
27.0455 | -97.8209
27.0492 | -97.8505
27.053 | -97.8913
27.0567 | -97.9229
27.0605 | -97.9537
27.0642 | -97.9863
27.068 | -98.023
27.0717 | -98.0617
27.0755 | -98.0996
27.0792 | -98.1344
27.083 | -98.1678
27.0867 | -98.2023
27.0905 | -98.2363
27.0942 | -98.2668
27.098 | -98.2996
27.1017 | -98.337
27.1055 | -98.3749
27.1092 | -98.4112
27.113 | -98.4476
27.1167 | -98.4859
27.1205 | -98.5243
27.1242 | -98.5612
27.128 | -98.5991
27.1317 | -98.6355
27.1355 | -98.6698
27.1392 | -98.7034
27.143 | -98.7375
27.1467 | -98.7697
27.1505 | -98.8
27.1542 | -98.8299
27.158 | -98.858
27.1617 | -98.8863
27.1655 | -98.9118
27.1692 | -98.9317
27.173 | -98.9553
27.1767 | -98.9751
27.1805 | -98.9907
27.1842 | -99.0099
27.188 | -99.0259
27.1917 | -99.037
27.1955 | -99.049
27.1992 | -99.0556
27.203 | -99.0587
27.2068 | -99.0591
27.2105 | -99.065
27.2143 | -99.0648
27.218 | -99.0619
27.2218 | -99.0612
27.2255 | -99.0593
27.2293 | -99.0524
27.233 | -99.0464
27.2368 | -99.0421
27.2405 | -99.0377
27.2443 | -99.0299
27.248 | -99.0179
27.2518 | -99.0064
27.2555 | -98.996
27.2593 | -98.9852
27.263 | -98.9762
27.2668 | -98.9708
27.2705 | -98.9641
27.2743 | -98.9549
27.278 | -98.9467
27.2818 | -98.9382
27.2855 | -98.9307
27.2893 | -98.923
27.293 | -98.9163
27.2968 | -98.9117
27.3005 | -98.9082
27.3043 | -98.902
27.308 | -98.896
27.3118 | -98.8899
27.3155 | -98.8861
27.3193 | -98.883
27.323 | -98.88
27.3268 | -98.8784
27.3305 | -98.8771
27.3343 | -98.8771
27.338 | -98.8786
27.3418 | -98.8801
27.3455 | -98.8812
27.3493 | -98.8814
27.353 | -98.8833
27.3568 | -98.8865
27.3605 | -98.8887
27.3643 | -98.8922
27.368 | -98.898
27.3718 | -98.9037
27.3755 | -98.9083
27.3793 | -98.9146
27.383 | -98.9196
27.3868 | -98.9222
27.3905 | -98.9222
27.3943 | -98.9207
27.398 | -98.9193
27.4018 | -98.9189
27.4055 | -98.92
27.4093 | -98.9195
27.413 | -98.9204
27.4168 | -98.9208
27.4205 | -98.9219
27.4243 | -98.9236
27.428 | -98.9229
27.4318 | -98.9224
27.4355 | -98.9225
27.4393 | -98.9216
27.443 | -98.9194
27.4468 | -98.9175
27.4505 | -98.916
27.4543 | -98.9138
27.458 | -98.9114
27.4618 | -98.91
27.4655 | -98.9085
27.4693 | -98.9071
27.473 | -98.9049
27.4768 | -98.9031
27.4805 | -98.9049
27.4843 | -98.9029
27.488 | -98.8952
27.4918 | -98.8889
27.4955 | -98.8839
27.4993 | -98.8772
27.503 | -98.869
27.5068 | -98.8616
27.5105 | -98.8555
27.5143 | -98.8478
27.518 | -98.8402
27.5218 | -98.8336
27.5256 | -98.8265
27.5293 | -98.8198
27.5331 | -98.8146
27.5368 | -98.8085
27.5406 | -98.8012
27.5443 | -98.7944
27.5481 | -98.7867
27.5518 | -98.7777
27.5556 | -98.7713
27.5593 | -98.7665
27.5631 | -98.7643
27.5668 | -98.7629
27.5706 | -98.7613
27.5743 | -98.7598
27.5781 | -98.7569
27.5818 | -98.7534
27.5856 | -98.7523
27.5893 | -98.7529
27.5931 | -98.7548
27.5968 | -98.754
27.6006 | -98.7541
27.6043 | -98.7552
27.6081 | -98.7552
27.6118 | -98.7536
27.6156 | -98.7538
27.6193 | -98.7558
27.6231 | -98.7571
27.6268 | -98.7574
27.6306 | -98.7588
27.6343 | -98.7618
27.6381 | -98.7663
27.6418 | -98.7702
27.6456 | -98.7744
27.6493 | -98.778
27.6531 | -98.7813
27.6568 | -98.7841
27.6606 | -98.788
27.6643 | -98.7936
27.6681 | -98.7974
27.6718 | -98.7995
27.6756 | -98.802
27.6793 | -98.806
27.6831 | -98.8089
27.6868 | -98.8131
27.6906 | -98.8182
27.6943 | -98.8248
27.6981 | -98.8322
27.7018 | -98.8385
27.7056 | -98.8446
27.7093 | -98.8512
27.7131 | -98.859
27.7168 | -98.8668
27.7206 | -98.8744
27.7243 | -98.8833
27.7281 | -98.8923
27.7318 | -98.9015
27.7356 | -98.9135
27.7393 | -98.9247
27.7431 | -98.9353
27.7468 | -98.9476
27.7506 | -98.9609
27.7543 | -98.9738
27.7581 | -98.9857
27.7618 | -98.998
27.7656 | -99.0118
27.7693 | -99.0248
27.7731 | -99.041
27.7768 | -99.0587
27.7806 | -99.0766
27.7843 | -99.0957
27.7881 | -99.1153
27.7918 | -99.1328
27.7956 | -99.1487
27.7993 | -99.1651
27.8031 | -99.1798
27.8068 | -99.1935
27.8106 | -99.208
27.8143 | -99.2232
27.8181 | -99.2393
27.8218 | -99.2556
27.8256 | -99.2713
27.8293 | -99.2871
27.8331 | -99.3039
27.8368 | -99.3208
27.8406 | -99.3382
27.8444 | -99.3525
27.8481 | -99.3679
27.8519 | -99.3825
27.8556 | -99.3999
27.8594 | -99.4172
27.8631 | -99.4344
27.8669 | -99.45
27.8706 | -99.465
27.8744 | -99.4818
27.8781 | -99.5005
27.8819 | -99.5186
27.8856 | -99.5363
27.8894 | -99.5551
27.8931 | -99.5754
27.8969 | -99.5959
27.9006 | -99.6153
27.9044 | -99.6348
27.9081 | -99.6555
27.9119 | -99.6745
27.9156 | -99.6931
27.9194 | -99.7119
27.9231 | -99.7284
27.9269 | -99.7431
27.9306 | -99.7586
27.9344 | -99.7777
27.9381 | -99.7979
27.9419 | -99.8169
27.9456 | -99.8317
27.9494 | -99.8417
27.9531 | -99.8509
27.9569 | -99.8607
27.9606 | -99.8705
27.9644 | -99.8794
27.9681 | -99.8908
27.9719 | -99.9006
27.9756 | -99.911
27.9794 | -99.9193
27.9831 | -99.924
27.9869 | -99.9311
27.9906 | -99.9375
27.9944 | -99.9396
27.9981 | -99.9416
28.0019 | -99.9447