## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 238 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 239
Hood | 02-92B/P | 178
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1010.6 | hPa |  | PRC PRO 01
Humidity | 52.9 | %/Hr
Internal temperature of calibration bench | 23.67 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1018 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 25.2 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.95 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.99 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3096 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 11/06/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 12/06/2018
Technical Manager | Anthony HUE | Approval |  | 13/06/2018
CEO | Robert GLEMOT | Approval |  | 13/06/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3881 | 57.3553
0.012209 | -35.9589 | 38.1282
0.018313 | -34.8994 | 28.0108
0.024417 | -34.4455 | 21.7619
0.030521 | -34.3406 | 18.0179
0.036626 | -34.2007 | 14.9443
0.04273 | -34.1087 | 12.5622
0.048834 | -34.0647 | 11.06
0.054938 | -33.9997 | 9.9267
0.061043 | -33.9383 | 8.8334
0.067147 | -33.9307 | 7.9955
0.073251 | -33.935 | 7.2197
0.079355 | -33.9567 | 6.5481
0.08546 | -33.9728 | 5.9498
0.091564 | -33.9691 | 5.5098
0.097668 | -33.9519 | 5.2638
0.10377 | -33.931 | 5.0127
0.10988 | -33.9148 | 4.7326
0.11598 | -33.9068 | 4.4374
0.12209 | -33.9102 | 4.1527
0.12819 | -33.9223 | 3.8684
0.13429 | -33.9354 | 3.5786
0.1404 | -33.9395 | 3.2961
0.1465 | -33.9421 | 3.0347
0.15261 | -33.9394 | 2.8109
0.15871 | -33.9328 | 2.6469
0.16482 | -33.9249 | 2.5431
0.17092 | -33.9237 | 2.5035
0.17702 | -33.9225 | 2.4297
0.18313 | -33.9226 | 2.3502
0.18923 | -33.9231 | 2.2387
0.19534 | -33.9225 | 2.1089
0.20144 | -33.9216 | 1.963
0.20754 | -33.9192 | 1.8391
0.21365 | -33.9188 | 1.722
0.21975 | -33.9191 | 1.6223
0.22586 | -33.9211 | 1.5343
0.23196 | -33.9246 | 1.4622
0.23807 | -33.9297 | 1.3998
0.24417 | -33.9347 | 1.34
0.25027 | -33.9397 | 1.2761
0.25638 | -33.9428 | 1.2119
0.26248 | -33.9454 | 1.1423
0.26859 | -33.947 | 1.0736
0.27469 | -33.9481 | 0.99871
0.2808 | -33.9499 | 0.92285
0.2869 | -33.9514 | 0.84986
0.293 | -33.9531 | 0.7813
0.29911 | -33.9547 | 0.72076
0.30521 | -33.9564 | 0.66146
0.31132 | -33.9578 | 0.62408
0.31742 | -33.9597 | 0.58196
0.32353 | -33.9607 | 0.54768
0.32963 | -33.9622 | 0.5195
0.33573 | -33.9639 | 0.49976
0.34184 | -33.9659 | 0.47863
0.34794 | -33.9687 | 0.45395
0.35405 | -33.9715 | 0.42924
0.36015 | -33.9747 | 0.40308
0.36626 | -33.9773 | 0.37439
0.37236 | -33.98 | 0.33611
0.37846 | -33.9828 | 0.28993
0.38457 | -33.984 | 0.23089
0.39067 | -33.9854 | 0.18265
0.39678 | -33.9864 | 0.12911
0.40288 | -33.9871 | 0.07332
0.40899 | -33.9873 | 0.01982
0.41509 | -33.9868 | -0.032434
0.42119 | -33.987 | -0.076958
0.4273 | -33.9872 | -0.12468
0.4334 | -33.9874 | -0.17242
0.43951 | -33.988 | -0.2176
0.44561 | -33.9892 | -0.25898
0.45172 | -33.9913 | -0.29242
0.45782 | -33.9926 | -0.33122
0.46392 | -33.9943 | -0.36922
0.47003 | -33.9959 | -0.40319
0.47613 | -33.9973 | -0.43896
0.48224 | -33.9983 | -0.47554
0.48834 | -33.9978 | -0.51228
0.49445 | -33.9981 | -0.54939
0.50055 | -33.998 | -0.58326
0.50665 | -33.9977 | -0.61946
0.51276 | -33.9971 | -0.6555
0.51886 | -33.9965 | -0.69122
0.52497 | -33.9956 | -0.71686
0.53107 | -33.9953 | -0.74813
0.53717 | -33.995 | -0.77786
0.54328 | -33.9945 | -0.80212
0.54938 | -33.9947 | -0.82595
0.55549 | -33.9951 | -0.84805
0.56159 | -33.9966 | -0.86031
0.5677 | -33.9972 | -0.87567
0.5738 | -33.9987 | -0.89293
0.5799 | -33.9999 | -0.90864
0.58601 | -34.0013 | -0.91919
0.59211 | -34.0034 | -0.92557
0.59822 | -34.0051 | -0.93538
0.60432 | -34.007 | -0.94797
0.61043 | -34.0086 | -0.95609
0.61653 | -34.0102 | -0.96657
0.62263 | -34.0121 | -0.97603
0.62874 | -34.013 | -0.98675
0.63484 | -34.0142 | -0.99611
0.64095 | -34.0149 | -1.0064
0.64705 | -34.0156 | -1.0163
0.65316 | -34.016 | -1.029
0.65926 | -34.0164 | -1.0454
0.66536 | -34.0155 | -1.0637
0.67147 | -34.0153 | -1.0775
0.67757 | -34.0152 | -1.0924
0.68368 | -34.0147 | -1.1122
0.68978 | -34.0142 | -1.1302
0.69589 | -34.0135 | -1.1501
0.70199 | -34.0128 | -1.1741
0.70809 | -34.0122 | -1.1941
0.7142 | -34.0118 | -1.2156
0.7203 | -34.0109 | -1.2359
0.72641 | -34.0106 | -1.2594
0.73251 | -34.0102 | -1.2805
0.73862 | -34.0108 | -1.3032
0.74472 | -34.0108 | -1.3241
0.75082 | -34.0108 | -1.3472
0.75693 | -34.0115 | -1.3701
0.76303 | -34.0124 | -1.393
0.76914 | -34.0145 | -1.4141
0.77524 | -34.0159 | -1.4346
0.78135 | -34.0175 | -1.4555
0.78745 | -34.0192 | -1.4762
0.79355 | -34.0208 | -1.4968
0.79966 | -34.0229 | -1.517
0.80576 | -34.0255 | -1.5374
0.81187 | -34.0274 | -1.56
0.81797 | -34.0297 | -1.5821
0.82408 | -34.0317 | -1.6034
0.83018 | -34.0338 | -1.6248
0.83628 | -34.036 | -1.6441
0.84239 | -34.0374 | -1.6657
0.84849 | -34.0392 | -1.6885
0.8546 | -34.0409 | -1.7093
0.8607 | -34.0424 | -1.7309
0.86681 | -34.0438 | -1.7508
0.87291 | -34.045 | -1.7706
0.87901 | -34.0458 | -1.7896
0.88512 | -34.0468 | -1.8087
0.89122 | -34.0478 | -1.8279
0.89733 | -34.0484 | -1.8444
0.90343 | -34.0491 | -1.8628
0.90953 | -34.0496 | -1.8779
0.91564 | -34.0498 | -1.8946
0.92174 | -34.0503 | -1.9103
0.92785 | -34.0507 | -1.9255
0.93395 | -34.051 | -1.94
0.94006 | -34.0515 | -1.9534
0.94616 | -34.0515 | -1.9656
0.95226 | -34.0518 | -1.9788
0.95837 | -34.0522 | -1.9931
0.96447 | -34.0527 | -2.0064
0.97058 | -34.0528 | -2.0202
0.97668 | -34.0528 | -2.034
0.98279 | -34.0528 | -2.0471
0.98889 | -34.0528 | -2.0581
0.99499 | -34.053 | -2.0719
1.0011 | -34.0531 | -2.0828
1.0072 | -34.053 | -2.0947
1.0133 | -34.0529 | -2.1057
1.0194 | -34.0527 | -2.1192
1.0255 | -34.0526 | -2.1333
1.0316 | -34.0522 | -2.1456
1.0377 | -34.052 | -2.1596
1.0438 | -34.0516 | -2.1728
1.0499 | -34.051 | -2.1886
1.056 | -34.0504 | -2.2033
1.0621 | -34.05 | -2.2181
1.0682 | -34.0497 | -2.231
1.0743 | -34.0492 | -2.247
1.0805 | -34.0489 | -2.262
1.0866 | -34.0481 | -2.2777
1.0927 | -34.0473 | -2.2927
1.0988 | -34.0469 | -2.3081
1.1049 | -34.0462 | -2.3229
1.111 | -34.0455 | -2.3389
1.1171 | -34.045 | -2.354
1.1232 | -34.0447 | -2.3699
1.1293 | -34.0444 | -2.3844
1.1354 | -34.0443 | -2.4001
1.1415 | -34.044 | -2.4163
1.1476 | -34.0436 | -2.4325
1.1537 | -34.0432 | -2.4478
1.1598 | -34.0435 | -2.4633
1.1659 | -34.0434 | -2.4772
1.172 | -34.0433 | -2.4917
1.1781 | -34.0432 | -2.5063
1.1842 | -34.0433 | -2.5212
1.1903 | -34.0438 | -2.5327
1.1964 | -34.0439 | -2.5466
1.2025 | -34.0442 | -2.5612
1.2086 | -34.0447 | -2.5742
1.2147 | -34.0449 | -2.5872
1.2209 | -34.0454 | -2.599
1.227 | -34.0464 | -2.6102
1.2331 | -34.047 | -2.6227
1.2392 | -34.0476 | -2.6351
1.2453 | -34.0482 | -2.6466
1.2514 | -34.0488 | -2.657
1.2575 | -34.0496 | -2.6684
1.2636 | -34.0507 | -2.6777
1.2697 | -34.0515 | -2.6873
1.2758 | -34.0524 | -2.6989
1.2819 | -34.0534 | -2.7093
1.288 | -34.0542 | -2.7205
1.2941 | -34.0553 | -2.7293
1.3002 | -34.0562 | -2.7376
1.3063 | -34.057 | -2.747
1.3124 | -34.0578 | -2.7563
1.3185 | -34.0588 | -2.7657
1.3246 | -34.0597 | -2.7752
1.3307 | -34.0608 | -2.7841
1.3368 | -34.0618 | -2.7931
1.3429 | -34.0627 | -2.8024
1.349 | -34.0634 | -2.812
1.3551 | -34.0643 | -2.8217
1.3613 | -34.0651 | -2.8303
1.3674 | -34.0659 | -2.8407
1.3735 | -34.0665 | -2.8507
1.3796 | -34.0672 | -2.8601
1.3857 | -34.0677 | -2.8696
1.3918 | -34.0684 | -2.8802
1.3979 | -34.0692 | -2.8896
1.404 | -34.0699 | -2.9014
1.4101 | -34.0705 | -2.9121
1.4162 | -34.071 | -2.9231
1.4223 | -34.0717 | -2.9349
1.4284 | -34.0724 | -2.9452
1.4345 | -34.0728 | -2.9571
1.4406 | -34.0733 | -2.9711
1.4467 | -34.0739 | -2.9832
1.4528 | -34.0745 | -2.9949
1.4589 | -34.0752 | -3.0072
1.465 | -34.0756 | -3.0184
1.4711 | -34.0763 | -3.0302
1.4772 | -34.077 | -3.044
1.4833 | -34.0776 | -3.0575
1.4894 | -34.0783 | -3.0701
1.4955 | -34.079 | -3.0825
1.5016 | -34.0798 | -3.0955
1.5078 | -34.0803 | -3.1099
1.5139 | -34.0809 | -3.1234
1.52 | -34.0815 | -3.1364
1.5261 | -34.082 | -3.1495
1.5322 | -34.0824 | -3.1641
1.5383 | -34.083 | -3.1771
1.5444 | -34.0837 | -3.1901
1.5505 | -34.0844 | -3.204
1.5566 | -34.085 | -3.2175
1.5627 | -34.0857 | -3.2318
1.5688 | -34.0865 | -3.2459
1.5749 | -34.087 | -3.259
1.581 | -34.0874 | -3.2728
1.5871 | -34.088 | -3.2871
1.5932 | -34.0885 | -3.3004
1.5993 | -34.0891 | -3.3152
1.6054 | -34.0896 | -3.3289
1.6115 | -34.0902 | -3.3435
1.6176 | -34.0907 | -3.3566
1.6237 | -34.091 | -3.3707
1.6298 | -34.0915 | -3.3837
1.6359 | -34.092 | -3.3969
1.642 | -34.0926 | -3.4092
1.6482 | -34.0929 | -3.4227
1.6543 | -34.0934 | -3.4368
1.6604 | -34.0939 | -3.4493
1.6665 | -34.0942 | -3.4621
1.6726 | -34.0945 | -3.4751
1.6787 | -34.095 | -3.4881
1.6848 | -34.0951 | -3.501
1.6909 | -34.0955 | -3.5142
1.697 | -34.0959 | -3.5268
1.7031 | -34.0961 | -3.5398
1.7092 | -34.0962 | -3.5523
1.7153 | -34.0965 | -3.5647
1.7214 | -34.0964 | -3.5773
1.7275 | -34.0965 | -3.5903
1.7336 | -34.0966 | -3.6036
1.7397 | -34.0969 | -3.6164
1.7458 | -34.097 | -3.6289
1.7519 | -34.0972 | -3.6414
1.758 | -34.0973 | -3.6537
1.7641 | -34.0976 | -3.6656
1.7702 | -34.0977 | -3.679
1.7763 | -34.0978 | -3.6907
1.7824 | -34.0979 | -3.7035
1.7885 | -34.098 | -3.7164
1.7947 | -34.0983 | -3.7287
1.8008 | -34.0984 | -3.7402
1.8069 | -34.0986 | -3.7521
1.813 | -34.0988 | -3.7646
1.8191 | -34.0989 | -3.7777
1.8252 | -34.0991 | -3.7908
1.8313 | -34.0994 | -3.8034
1.8374 | -34.0997 | -3.8156
1.8435 | -34.0999 | -3.8279
1.8496 | -34.1001 | -3.8413
1.8557 | -34.1004 | -3.8544
1.8618 | -34.1008 | -3.867
1.8679 | -34.1011 | -3.8794
1.874 | -34.1013 | -3.8917
1.8801 | -34.1016 | -3.9045
1.8862 | -34.102 | -3.9168
1.8923 | -34.1023 | -3.9293
1.8984 | -34.1025 | -3.9417
1.9045 | -34.1027 | -3.9536
1.9106 | -34.1029 | -3.9662
1.9167 | -34.1031 | -3.9788
1.9228 | -34.1034 | -3.9909
1.9289 | -34.1036 | -4.0032
1.9351 | -34.1039 | -4.0144
1.9412 | -34.1043 | -4.026
1.9473 | -34.1047 | -4.0384
1.9534 | -34.105 | -4.0502
1.9595 | -34.1054 | -4.0623
1.9656 | -34.1058 | -4.0751
1.9717 | -34.1061 | -4.0871
1.9778 | -34.1066 | -4.0989
1.9839 | -34.107 | -4.1108
1.99 | -34.1073 | -4.1222
1.9961 | -34.1075 | -4.1347
2.0022 | -34.1079 | -4.1466
2.0083 | -34.1081 | -4.1581
2.0144 | -34.1085 | -4.1698
2.0205 | -34.1089 | -4.182
2.0266 | -34.1094 | -4.1938
2.0327 | -34.1098 | -4.2058
2.0388 | -34.1103 | -4.2177
2.0449 | -34.1106 | -4.2295
2.051 | -34.111 | -4.2413
2.0571 | -34.1115 | -4.253
2.0632 | -34.1118 | -4.2649
2.0693 | -34.1121 | -4.2775
2.0754 | -34.1125 | -4.2892
2.0816 | -34.113 | -4.3018
2.0877 | -34.1133 | -4.3138
2.0938 | -34.1138 | -4.3263
2.0999 | -34.1141 | -4.3384
2.106 | -34.1145 | -4.3505
2.1121 | -34.1148 | -4.3629
2.1182 | -34.1152 | -4.374
2.1243 | -34.1156 | -4.3852
2.1304 | -34.1159 | -4.3967
2.1365 | -34.1164 | -4.4089
2.1426 | -34.1167 | -4.4213
2.1487 | -34.1171 | -4.433
2.1548 | -34.1177 | -4.4452
2.1609 | -34.1183 | -4.4571
2.167 | -34.1188 | -4.4687
2.1731 | -34.1192 | -4.4812
2.1792 | -34.1196 | -4.4926
2.1853 | -34.1201 | -4.505
2.1914 | -34.1205 | -4.5168
2.1975 | -34.1211 | -4.5288
2.2036 | -34.1214 | -4.5414
2.2097 | -34.1218 | -4.5533
2.2158 | -34.1223 | -4.5643
2.222 | -34.1227 | -4.5752
2.2281 | -34.1231 | -4.587
2.2342 | -34.1235 | -4.5988
2.2403 | -34.1239 | -4.6098
2.2464 | -34.1243 | -4.6209
2.2525 | -34.1247 | -4.6328
2.2586 | -34.1251 | -4.6438
2.2647 | -34.1256 | -4.6545
2.2708 | -34.126 | -4.6662
2.2769 | -34.1265 | -4.6773
2.283 | -34.1268 | -4.6888
2.2891 | -34.1272 | -4.7002
2.2952 | -34.1275 | -4.7111
2.3013 | -34.1278 | -4.7224
2.3074 | -34.1282 | -4.7332
2.3135 | -34.1286 | -4.745
2.3196 | -34.1289 | -4.7555
2.3257 | -34.1293 | -4.7666
2.3318 | -34.1295 | -4.7781
2.3379 | -34.1298 | -4.7893
2.344 | -34.1301 | -4.8011
2.3501 | -34.1304 | -4.8121
2.3562 | -34.1307 | -4.8228
2.3623 | -34.131 | -4.8336
2.3685 | -34.1312 | -4.8445
2.3746 | -34.1315 | -4.8561
2.3807 | -34.1318 | -4.8669
2.3868 | -34.1321 | -4.878
2.3929 | -34.1324 | -4.889
2.399 | -34.1326 | -4.8998
2.4051 | -34.133 | -4.9107
2.4112 | -34.1334 | -4.9217
2.4173 | -34.1337 | -4.9327
2.4234 | -34.1338 | -4.9438
2.4295 | -34.1341 | -4.955
2.4356 | -34.1344 | -4.9662
2.4417 | -34.1347 | -4.9777
2.4478 | -34.1349 | -4.9889
2.4539 | -34.1352 | -4.9995
2.46 | -34.1354 | -5.0113
2.4661 | -34.1357 | -5.0227
2.4722 | -34.136 | -5.0332
2.4783 | -34.1362 | -5.0439
2.4844 | -34.1364 | -5.0553
2.4905 | -34.1366 | -5.0666
2.4966 | -34.1369 | -5.0777
2.5027 | -34.1371 | -5.0885
2.5089 | -34.1373 | -5.0996
2.515 | -34.1374 | -5.1108
2.5211 | -34.1376 | -5.1218
2.5272 | -34.1378 | -5.133
2.5333 | -34.1381 | -5.1443
2.5394 | -34.1384 | -5.1554
2.5455 | -34.1386 | -5.1671
2.5516 | -34.1388 | -5.1777
2.5577 | -34.1391 | -5.1894
2.5638 | -34.1394 | -5.2015
2.5699 | -34.1395 | -5.2132
2.576 | -34.1397 | -5.2251
2.5821 | -34.1399 | -5.2367
2.5882 | -34.1402 | -5.2481
2.5943 | -34.1404 | -5.2592
2.6004 | -34.1407 | -5.2704
2.6065 | -34.1409 | -5.2811
2.6126 | -34.1411 | -5.2918
2.6187 | -34.1414 | -5.3026
2.6248 | -34.1416 | -5.3136
2.6309 | -34.1419 | -5.3253
2.637 | -34.1422 | -5.3373
2.6431 | -34.1425 | -5.348
2.6492 | -34.1427 | -5.3596
2.6554 | -34.143 | -5.3708
2.6615 | -34.1432 | -5.382
2.6676 | -34.1436 | -5.3938
2.6737 | -34.1439 | -5.4052
2.6798 | -34.1443 | -5.4166
2.6859 | -34.1446 | -5.4282
2.692 | -34.1449 | -5.4393
2.6981 | -34.1452 | -5.4508
2.7042 | -34.1455 | -5.4622
2.7103 | -34.1458 | -5.4731
2.7164 | -34.1462 | -5.4843
2.7225 | -34.1465 | -5.496
2.7286 | -34.1469 | -5.5073
2.7347 | -34.1472 | -5.5181
2.7408 | -34.1476 | -5.5284
2.7469 | -34.1479 | -5.54
2.753 | -34.1483 | -5.5509
2.7591 | -34.1487 | -5.5625
2.7652 | -34.149 | -5.5742
2.7713 | -34.1492 | -5.5852
2.7774 | -34.1496 | -5.5965
2.7835 | -34.1499 | -5.6077
2.7896 | -34.1502 | -5.6185
2.7958 | -34.1506 | -5.6296
2.8019 | -34.151 | -5.6402
2.808 | -34.1513 | -5.6513
2.8141 | -34.1518 | -5.6627
2.8202 | -34.1522 | -5.6739
2.8263 | -34.1526 | -5.6851
2.8324 | -34.1531 | -5.6967
2.8385 | -34.1536 | -5.7079
2.8446 | -34.154 | -5.7188
2.8507 | -34.1545 | -5.73
2.8568 | -34.155 | -5.7407
2.8629 | -34.1553 | -5.7516
2.869 | -34.1557 | -5.763
2.8751 | -34.1562 | -5.774
2.8812 | -34.1567 | -5.7856
2.8873 | -34.1571 | -5.7968
2.8934 | -34.1575 | -5.8081
2.8995 | -34.158 | -5.8195
2.9056 | -34.1583 | -5.831
2.9117 | -34.1588 | -5.8423
2.9178 | -34.1593 | -5.8533
2.9239 | -34.1598 | -5.8652
2.93 | -34.1603 | -5.877
2.9361 | -34.1607 | -5.8881
2.9423 | -34.1611 | -5.899
2.9484 | -34.1617 | -5.9104
2.9545 | -34.1622 | -5.9217
2.9606 | -34.1627 | -5.9334
2.9667 | -34.1632 | -5.9449
2.9728 | -34.1636 | -5.9565
2.9789 | -34.1642 | -5.9679
2.985 | -34.1647 | -5.9795
2.9911 | -34.1652 | -5.9912
2.9972 | -34.1657 | -6.0026
3.0033 | -34.1663 | -6.0144
3.0094 | -34.1667 | -6.0258
3.0155 | -34.1673 | -6.0371
3.0216 | -34.1678 | -6.0487
3.0277 | -34.1682 | -6.0607
3.0338 | -34.1687 | -6.0724
3.0399 | -34.1693 | -6.084
3.046 | -34.1697 | -6.0959
3.0521 | -34.1701 | -6.1074
3.0582 | -34.1705 | -6.1198
3.0643 | -34.171 | -6.1316
3.0704 | -34.1714 | -6.1438
3.0765 | -34.1718 | -6.1552
3.0827 | -34.1722 | -6.1668
3.0888 | -34.1727 | -6.1784
3.0949 | -34.1732 | -6.1901
3.101 | -34.1736 | -6.2019
3.1071 | -34.1741 | -6.214
3.1132 | -34.1745 | -6.2261
3.1193 | -34.175 | -6.238
3.1254 | -34.1755 | -6.2496
3.1315 | -34.176 | -6.2619
3.1376 | -34.1764 | -6.2737
3.1437 | -34.1768 | -6.2856
3.1498 | -34.1773 | -6.298
3.1559 | -34.1777 | -6.3098
3.162 | -34.1781 | -6.3223
3.1681 | -34.1784 | -6.3346
3.1742 | -34.1789 | -6.3465
3.1803 | -34.1793 | -6.3585
3.1864 | -34.1798 | -6.3701
3.1925 | -34.1802 | -6.382
3.1986 | -34.1806 | -6.3944
3.2047 | -34.181 | -6.4068
3.2108 | -34.1813 | -6.4183
3.2169 | -34.1817 | -6.4304
3.223 | -34.1821 | -6.4424
3.2292 | -34.1824 | -6.4546
3.2353 | -34.1827 | -6.4658
3.2414 | -34.1831 | -6.4776
3.2475 | -34.1834 | -6.4893
3.2536 | -34.1838 | -6.5011
3.2597 | -34.1842 | -6.5124
3.2658 | -34.1845 | -6.5236
3.2719 | -34.1848 | -6.535
3.278 | -34.1852 | -6.5469
3.2841 | -34.1855 | -6.5583
3.2902 | -34.1859 | -6.5701
3.2963 | -34.1862 | -6.582
3.3024 | -34.1864 | -6.5933
3.3085 | -34.1868 | -6.6052
3.3146 | -34.1871 | -6.6167
3.3207 | -34.1875 | -6.6285
3.3268 | -34.1878 | -6.6403
3.3329 | -34.1881 | -6.6514
3.339 | -34.1883 | -6.6625
3.3451 | -34.1886 | -6.6741
3.3512 | -34.1889 | -6.6851
3.3573 | -34.1892 | -6.6966
3.3634 | -34.1895 | -6.7081
3.3696 | -34.1898 | -6.7198
3.3757 | -34.19 | -6.7309
3.3818 | -34.1904 | -6.7421
3.3879 | -34.1906 | -6.7535
3.394 | -34.1909 | -6.7651
3.4001 | -34.1912 | -6.7763
3.4062 | -34.1914 | -6.7875
3.4123 | -34.1916 | -6.7984
3.4184 | -34.1919 | -6.81
3.4245 | -34.1922 | -6.8211
3.4306 | -34.1924 | -6.8321
3.4367 | -34.1927 | -6.8429
3.4428 | -34.193 | -6.8535
3.4489 | -34.1932 | -6.8648
3.455 | -34.1934 | -6.8759
3.4611 | -34.1937 | -6.8875
3.4672 | -34.1939 | -6.8988
3.4733 | -34.1942 | -6.9102
3.4794 | -34.1945 | -6.9213
3.4855 | -34.1947 | -6.932
3.4916 | -34.195 | -6.9426
3.4977 | -34.1952 | -6.9542
3.5038 | -34.1955 | -6.9654
3.5099 | -34.1957 | -6.976
3.5161 | -34.1959 | -6.9866
3.5222 | -34.1961 | -6.9972
3.5283 | -34.1963 | -7.0083
3.5344 | -34.1965 | -7.0194
3.5405 | -34.1968 | -7.0302
3.5466 | -34.1971 | -7.0414
3.5527 | -34.1973 | -7.0525
3.5588 | -34.1975 | -7.0633
3.5649 | -34.1976 | -7.0743
3.571 | -34.1978 | -7.0852
3.5771 | -34.1981 | -7.0966
3.5832 | -34.1983 | -7.1067
3.5893 | -34.1986 | -7.1176
3.5954 | -34.1989 | -7.1288
3.6015 | -34.1991 | -7.1399
3.6076 | -34.1994 | -7.1508
3.6137 | -34.1996 | -7.1611
3.6198 | -34.2 | -7.1716
3.6259 | -34.2002 | -7.1828
3.632 | -34.2005 | -7.1937
3.6381 | -34.2008 | -7.2047
3.6442 | -34.201 | -7.2156
3.6503 | -34.2013 | -7.226
3.6565 | -34.2016 | -7.237
3.6626 | -34.2018 | -7.2479
3.6687 | -34.2021 | -7.2589
3.6748 | -34.2024 | -7.2696
3.6809 | -34.2028 | -7.2805
3.687 | -34.203 | -7.2917
3.6931 | -34.2034 | -7.3022
3.6992 | -34.2037 | -7.3131
3.7053 | -34.204 | -7.3238
3.7114 | -34.2042 | -7.3348
3.7175 | -34.2045 | -7.3458
3.7236 | -34.2049 | -7.3561
3.7297 | -34.2051 | -7.3671
3.7358 | -34.2054 | -7.3783
3.7419 | -34.2057 | -7.3894
3.748 | -34.206 | -7.4003
3.7541 | -34.2063 | -7.4112
3.7602 | -34.2067 | -7.4219
3.7663 | -34.207 | -7.4326
3.7724 | -34.2073 | -7.4436
3.7785 | -34.2077 | -7.4542
3.7846 | -34.208 | -7.4649
3.7907 | -34.2084 | -7.4761
3.7969 | -34.2087 | -7.4871
3.803 | -34.209 | -7.4976
3.8091 | -34.2094 | -7.5084
3.8152 | -34.2097 | -7.5191
3.8213 | -34.21 | -7.5302
3.8274 | -34.2104 | -7.5408
3.8335 | -34.2108 | -7.5518
3.8396 | -34.2112 | -7.5627
3.8457 | -34.2116 | -7.5737
3.8518 | -34.212 | -7.5845
3.8579 | -34.2125 | -7.5953
3.864 | -34.2128 | -7.6064
3.8701 | -34.2132 | -7.6177
3.8762 | -34.2136 | -7.6284
3.8823 | -34.2139 | -7.6394
3.8884 | -34.2143 | -7.6504
3.8945 | -34.2148 | -7.6616
3.9006 | -34.2151 | -7.6729
3.9067 | -34.2155 | -7.6842
3.9128 | -34.2159 | -7.6953
3.9189 | -34.2163 | -7.7059
3.925 | -34.2166 | -7.717
3.9311 | -34.217 | -7.7275
3.9372 | -34.2174 | -7.7386
3.9434 | -34.2177 | -7.7496
3.9495 | -34.2182 | -7.7607
3.9556 | -34.2186 | -7.7714
3.9617 | -34.219 | -7.782
3.9678 | -34.2194 | -7.7926
3.9739 | -34.2198 | -7.8036
3.98 | -34.2201 | -7.8146
3.9861 | -34.2205 | -7.8255
3.9922 | -34.2208 | -7.8364
3.9983 | -34.2212 | -7.8469
4.0044 | -34.2216 | -7.8574
4.0105 | -34.222 | -7.8683
4.0166 | -34.2224 | -7.8793
4.0227 | -34.2228 | -7.8902
4.0288 | -34.2232 | -7.9016
4.0349 | -34.2235 | -7.9123
4.041 | -34.2238 | -7.9231
4.0471 | -34.2241 | -7.9341
4.0532 | -34.2245 | -7.9447
4.0593 | -34.225 | -7.9558
4.0654 | -34.2253 | -7.9664
4.0715 | -34.2257 | -7.9778
4.0776 | -34.2261 | -7.9886
4.0838 | -34.2265 | -7.9999
4.0899 | -34.2269 | -8.0105
4.096 | -34.2272 | -8.0213
4.1021 | -34.2276 | -8.032
4.1082 | -34.2279 | -8.0426
4.1143 | -34.2283 | -8.0532
4.1204 | -34.2287 | -8.0637
4.1265 | -34.229 | -8.0742
4.1326 | -34.2294 | -8.0847
4.1387 | -34.2298 | -8.0954
4.1448 | -34.2302 | -8.1063
4.1509 | -34.2305 | -8.1171
4.157 | -34.2308 | -8.1284
4.1631 | -34.2312 | -8.1394
4.1692 | -34.2315 | -8.15
4.1753 | -34.2319 | -8.1607
4.1814 | -34.2322 | -8.1711
4.1875 | -34.2326 | -8.1822
4.1936 | -34.2329 | -8.1931
4.1997 | -34.2333 | -8.2035
4.2058 | -34.2336 | -8.2146
4.2119 | -34.2339 | -8.2252
4.218 | -34.2342 | -8.2356
4.2241 | -34.2346 | -8.2461
4.2303 | -34.2349 | -8.2567
4.2364 | -34.2353 | -8.2674
4.2425 | -34.2355 | -8.2782
4.2486 | -34.2358 | -8.2891
4.2547 | -34.2361 | -8.2995
4.2608 | -34.2364 | -8.3103
4.2669 | -34.2366 | -8.3206
4.273 | -34.237 | -8.3315
4.2791 | -34.2373 | -8.3422
4.2852 | -34.2376 | -8.353
4.2913 | -34.2379 | -8.3634
4.2974 | -34.2382 | -8.3743
4.3035 | -34.2385 | -8.3853
4.3096 | -34.2388 | -8.3962
4.3157 | -34.2391 | -8.4066
4.3218 | -34.2394 | -8.4172
4.3279 | -34.2397 | -8.4279
4.334 | -34.24 | -8.4385
4.3401 | -34.2403 | -8.4486
4.3462 | -34.2406 | -8.4594
4.3523 | -34.241 | -8.4696
4.3584 | -34.2412 | -8.4797
4.3645 | -34.2415 | -8.4902
4.3707 | -34.2418 | -8.5005
4.3768 | -34.2421 | -8.5111
4.3829 | -34.2424 | -8.522
4.389 | -34.2427 | -8.5326
4.3951 | -34.243 | -8.5429
4.4012 | -34.2433 | -8.5538
4.4073 | -34.2435 | -8.5642
4.4134 | -34.2438 | -8.5746
4.4195 | -34.2442 | -8.5851
4.4256 | -34.2445 | -8.5955
4.4317 | -34.2448 | -8.6059
4.4378 | -34.2451 | -8.6161
4.4439 | -34.2454 | -8.6263
4.45 | -34.2456 | -8.6368
4.4561 | -34.2458 | -8.6473
4.4622 | -34.2461 | -8.6579
4.4683 | -34.2464 | -8.6683
4.4744 | -34.2467 | -8.6791
4.4805 | -34.247 | -8.6892
4.4866 | -34.2473 | -8.6996
4.4927 | -34.2476 | -8.7104
4.4988 | -34.2478 | -8.7208
4.5049 | -34.2481 | -8.7309
4.511 | -34.2485 | -8.741
4.5172 | -34.2488 | -8.7514
4.5233 | -34.249 | -8.7623
4.5294 | -34.2493 | -8.7726
4.5355 | -34.2496 | -8.7831
4.5416 | -34.2499 | -8.7934
4.5477 | -34.2502 | -8.8039
4.5538 | -34.2504 | -8.8139
4.5599 | -34.2508 | -8.8243
4.566 | -34.2511 | -8.8347
4.5721 | -34.2513 | -8.8451
4.5782 | -34.2517 | -8.8558
4.5843 | -34.252 | -8.8663
4.5904 | -34.2523 | -8.8768
4.5965 | -34.2526 | -8.8871
4.6026 | -34.2529 | -8.898
4.6087 | -34.2532 | -8.9083
4.6148 | -34.2534 | -8.9191
4.6209 | -34.2538 | -8.9295
4.627 | -34.2541 | -8.9401
4.6331 | -34.2543 | -8.9504
4.6392 | -34.2546 | -8.961
4.6453 | -34.2549 | -8.9714
4.6514 | -34.2552 | -8.9819
4.6576 | -34.2555 | -8.9927
4.6637 | -34.2558 | -9.0028
4.6698 | -34.256 | -9.013
4.6759 | -34.2563 | -9.0239
4.682 | -34.2566 | -9.0342
4.6881 | -34.2568 | -9.0446
4.6942 | -34.2571 | -9.055
4.7003 | -34.2574 | -9.0655
4.7064 | -34.2577 | -9.0757
4.7125 | -34.2581 | -9.0865
4.7186 | -34.2583 | -9.0967
4.7247 | -34.2586 | -9.1072
4.7308 | -34.259 | -9.1176
4.7369 | -34.2592 | -9.1281
4.743 | -34.2595 | -9.139
4.7491 | -34.2598 | -9.1498
4.7552 | -34.2601 | -9.1603
4.7613 | -34.2604 | -9.1709
4.7674 | -34.2606 | -9.1813
4.7735 | -34.2609 | -9.1922
4.7796 | -34.2613 | -9.203
4.7857 | -34.2616 | -9.2136
4.7918 | -34.2619 | -9.2239
4.7979 | -34.2622 | -9.2343
4.8041 | -34.2625 | -9.2448
4.8102 | -34.2628 | -9.255
4.8163 | -34.2631 | -9.2654
4.8224 | -34.2634 | -9.276
4.8285 | -34.2637 | -9.2862
4.8346 | -34.2641 | -9.2965
4.8407 | -34.2643 | -9.307
4.8468 | -34.2647 | -9.3178
4.8529 | -34.265 | -9.3286
4.859 | -34.2653 | -9.3395
4.8651 | -34.2656 | -9.3497
4.8712 | -34.2659 | -9.36
4.8773 | -34.2662 | -9.3706
4.8834 | -34.2665 | -9.3811
4.8895 | -34.2668 | -9.3914
4.8956 | -34.2671 | -9.4019
4.9017 | -34.2673 | -9.4122
4.9078 | -34.2677 | -9.423
4.9139 | -34.268 | -9.4331
4.92 | -34.2683 | -9.4433
4.9261 | -34.2687 | -9.4535
4.9322 | -34.269 | -9.464
4.9383 | -34.2693 | -9.4745
4.9445 | -34.2697 | -9.485
4.9506 | -34.2699 | -9.4953
4.9567 | -34.2703 | -9.5055
4.9628 | -34.2706 | -9.516
4.9689 | -34.2709 | -9.5268
4.975 | -34.2712 | -9.5376
4.9811 | -34.2715 | -9.5478
4.9872 | -34.2719 | -9.5582
4.9933 | -34.2723 | -9.569
4.9994 | -34.2726 | -9.5799
5.0055 | -34.2729 | -9.5905
5.0116 | -34.2733 | -9.6014
5.0177 | -34.2736 | -9.6122
5.0238 | -34.274 | -9.6231
5.0299 | -34.2744 | -9.634
5.036 | -34.2747 | -9.6445
5.0421 | -34.275 | -9.655
5.0482 | -34.2754 | -9.6655
5.0543 | -34.2757 | -9.6758
5.0604 | -34.2761 | -9.6863
5.0665 | -34.2764 | -9.6966
5.0726 | -34.2768 | -9.7075
5.0787 | -34.2771 | -9.7183
5.0848 | -34.2775 | -9.7289
5.091 | -34.2778 | -9.7392
5.0971 | -34.2782 | -9.7494
5.1032 | -34.2785 | -9.7603
5.1093 | -34.2789 | -9.7709
5.1154 | -34.2792 | -9.7812
5.1215 | -34.2795 | -9.7915
5.1276 | -34.2798 | -9.8018
5.1337 | -34.2803 | -9.812
5.1398 | -34.2806 | -9.8223
5.1459 | -34.2809 | -9.8331
5.152 | -34.2813 | -9.8438
5.1581 | -34.2816 | -9.8546
5.1642 | -34.282 | -9.8652
5.1703 | -34.2823 | -9.8758
5.1764 | -34.2827 | -9.8862
5.1825 | -34.2831 | -9.8964
5.1886 | -34.2834 | -9.9067
5.1947 | -34.2837 | -9.9172
5.2008 | -34.2841 | -9.9279
5.2069 | -34.2845 | -9.9387
5.213 | -34.2848 | -9.949
5.2191 | -34.2852 | -9.9596
5.2252 | -34.2855 | -9.9703
5.2314 | -34.2859 | -9.981
5.2375 | -34.2862 | -9.9913
5.2436 | -34.2865 | -10.0019
5.2497 | -34.2869 | -10.0122
5.2558 | -34.2873 | -10.0231
5.2619 | -34.2877 | -10.034
5.268 | -34.288 | -10.0447
5.2741 | -34.2884 | -10.0553
5.2802 | -34.2888 | -10.0663
5.2863 | -34.2892 | -10.0768
5.2924 | -34.2895 | -10.0876
5.2985 | -34.2899 | -10.0983
5.3046 | -34.2903 | -10.1089
5.3107 | -34.2907 | -10.1198
5.3168 | -34.291 | -10.1304
5.3229 | -34.2914 | -10.1411
5.329 | -34.2917 | -10.1517
5.3351 | -34.2921 | -10.162
5.3412 | -34.2925 | -10.1726
5.3473 | -34.2928 | -10.1832
5.3534 | -34.2932 | -10.194
5.3595 | -34.2936 | -10.2047
5.3656 | -34.294 | -10.2153
5.3717 | -34.2943 | -10.2263
5.3779 | -34.2947 | -10.2372
5.384 | -34.2951 | -10.2478
5.3901 | -34.2955 | -10.2588
5.3962 | -34.2958 | -10.2692
5.4023 | -34.2962 | -10.2801
5.4084 | -34.2966 | -10.2907
5.4145 | -34.297 | -10.3011
5.4206 | -34.2974 | -10.3118
5.4267 | -34.2977 | -10.3228
5.4328 | -34.2981 | -10.3335
5.4389 | -34.2985 | -10.3439
5.445 | -34.2988 | -10.3549
5.4511 | -34.2992 | -10.3653
5.4572 | -34.2996 | -10.3761
5.4633 | -34.3 | -10.3867
5.4694 | -34.3003 | -10.3974
5.4755 | -34.3007 | -10.4078
5.4816 | -34.3011 | -10.4185
5.4877 | -34.3015 | -10.4291
5.4938 | -34.3018 | -10.4395
5.4999 | -34.3022 | -10.45
5.506 | -34.3026 | -10.4607
5.5121 | -34.303 | -10.4712
5.5183 | -34.3034 | -10.4817
5.5244 | -34.3037 | -10.4926
5.5305 | -34.3041 | -10.5031
5.5366 | -34.3045 | -10.5135
5.5427 | -34.3048 | -10.5243
5.5488 | -34.3052 | -10.535
5.5549 | -34.3055 | -10.5457
5.561 | -34.3059 | -10.5566
5.5671 | -34.3063 | -10.5671
5.5732 | -34.3067 | -10.578
5.5793 | -34.3071 | -10.5891
5.5854 | -34.3075 | -10.5997
5.5915 | -34.3078 | -10.6107
5.5976 | -34.3082 | -10.6214
5.6037 | -34.3086 | -10.6322
5.6098 | -34.3089 | -10.6428
5.6159 | -34.3093 | -10.6533
5.622 | -34.3097 | -10.6641
5.6281 | -34.3101 | -10.6748
5.6342 | -34.3104 | -10.6859
5.6403 | -34.3108 | -10.6966
5.6464 | -34.3112 | -10.7074
5.6525 | -34.3115 | -10.7182
5.6586 | -34.3119 | -10.7286
5.6648 | -34.3124 | -10.7393
5.6709 | -34.3128 | -10.7501
5.677 | -34.3131 | -10.7609
5.6831 | -34.3135 | -10.7717
5.6892 | -34.3139 | -10.7826
5.6953 | -34.3142 | -10.7933
5.7014 | -34.3146 | -10.8041
5.7075 | -34.315 | -10.8147
5.7136 | -34.3153 | -10.8254
5.7197 | -34.3157 | -10.8361
5.7258 | -34.316 | -10.8465
5.7319 | -34.3164 | -10.8573
5.738 | -34.3168 | -10.868
5.7441 | -34.3173 | -10.8786
5.7502 | -34.3177 | -10.889
5.7563 | -34.318 | -10.8997
5.7624 | -34.3184 | -10.9106
5.7685 | -34.3188 | -10.9214
5.7746 | -34.3192 | -10.9323
5.7807 | -34.3195 | -10.9431
5.7868 | -34.3198 | -10.954
5.7929 | -34.3202 | -10.9646
5.799 | -34.3206 | -10.9751
5.8052 | -34.3209 | -10.9856
5.8113 | -34.3213 | -10.9962
5.8174 | -34.3217 | -11.0068
5.8235 | -34.3221 | -11.0174
5.8296 | -34.3224 | -11.0283
5.8357 | -34.3228 | -11.0389
5.8418 | -34.3231 | -11.0498
5.8479 | -34.3235 | -11.0605
5.854 | -34.3239 | -11.0713
5.8601 | -34.3243 | -11.0819
5.8662 | -34.3246 | -11.0927
5.8723 | -34.325 | -11.1032
5.8784 | -34.3253 | -11.1138
5.8845 | -34.3257 | -11.1243
5.8906 | -34.326 | -11.1351
5.8967 | -34.3264 | -11.1456
5.9028 | -34.3268 | -11.1561
5.9089 | -34.3273 | -11.1666
5.915 | -34.3276 | -11.177
5.9211 | -34.328 | -11.1876
5.9272 | -34.3283 | -11.198
5.9333 | -34.3287 | -11.2085
5.9394 | -34.3291 | -11.2192
5.9455 | -34.3295 | -11.2296
5.9517 | -34.3299 | -11.24
5.9578 | -34.3303 | -11.2508
5.9639 | -34.3307 | -11.261
5.97 | -34.3311 | -11.2717
5.9761 | -34.3315 | -11.2823
5.9822 | -34.3319 | -11.2929
5.9883 | -34.3322 | -11.3035
5.9944 | -34.3326 | -11.3143
6.0005 | -34.333 | -11.325
6.0066 | -34.3334 | -11.3354
6.0127 | -34.3338 | -11.3457
6.0188 | -34.3342 | -11.3566
6.0249 | -34.3345 | -11.3673
6.031 | -34.3349 | -11.3777
6.0371 | -34.3353 | -11.3883
6.0432 | -34.3357 | -11.3988
6.0493 | -34.336 | -11.4091
6.0554 | -34.3365 | -11.4195
6.0615 | -34.3369 | -11.4301
6.0676 | -34.3373 | -11.4409
6.0737 | -34.3377 | -11.4516
6.0798 | -34.338 | -11.4621
6.0859 | -34.3384 | -11.4726
6.0921 | -34.3388 | -11.4829
6.0982 | -34.3392 | -11.4935
6.1043 | -34.3396 | -11.5039
6.1104 | -34.34 | -11.5147
6.1165 | -34.3404 | -11.5247
6.1226 | -34.3408 | -11.5352
6.1287 | -34.3412 | -11.5457
6.1348 | -34.3416 | -11.5562
6.1409 | -34.342 | -11.5667
6.147 | -34.3423 | -11.5771
6.1531 | -34.3427 | -11.5875
6.1592 | -34.3431 | -11.5979
6.1653 | -34.3434 | -11.6083
6.1714 | -34.3438 | -11.6188
6.1775 | -34.3442 | -11.6294
6.1836 | -34.3446 | -11.64
6.1897 | -34.345 | -11.6504
6.1958 | -34.3454 | -11.6612
6.2019 | -34.3458 | -11.6718
6.208 | -34.3462 | -11.6823
6.2141 | -34.3466 | -11.6927
6.2202 | -34.347 | -11.7033
6.2263 | -34.3474 | -11.7137
6.2325 | -34.3477 | -11.7238
6.2386 | -34.3481 | -11.734
6.2447 | -34.3485 | -11.7448
6.2508 | -34.3489 | -11.7552
6.2569 | -34.3494 | -11.7656
6.263 | -34.3497 | -11.7761
6.2691 | -34.3501 | -11.7864
6.2752 | -34.3505 | -11.797
6.2813 | -34.3508 | -11.8074
6.2874 | -34.3512 | -11.8177
6.2935 | -34.3516 | -11.8281
6.2996 | -34.3519 | -11.8385
6.3057 | -34.3523 | -11.849
6.3118 | -34.3527 | -11.8595
6.3179 | -34.353 | -11.8699
6.324 | -34.3534 | -11.8803
6.3301 | -34.3538 | -11.8907
6.3362 | -34.3542 | -11.901
6.3423 | -34.3546 | -11.9116
6.3484 | -34.355 | -11.9221
6.3545 | -34.3554 | -11.9321
6.3606 | -34.3558 | -11.9423
6.3667 | -34.3561 | -11.9525
6.3728 | -34.3565 | -11.963
6.379 | -34.3569 | -11.9734
6.3851 | -34.3573 | -11.9835
6.3912 | -34.3576 | -11.9938
6.3973 | -34.358 | -12.0042
6.4034 | -34.3584 | -12.0144
6.4095 | -34.3587 | -12.0248
6.4156 | -34.3591 | -12.0352
6.4217 | -34.3595 | -12.0455
6.4278 | -34.3599 | -12.056
6.4339 | -34.3603 | -12.0662
6.44 | -34.3607 | -12.0765
6.4461 | -34.361 | -12.0871
6.4522 | -34.3614 | -12.0972
6.4583 | -34.3617 | -12.1077
6.4644 | -34.3621 | -12.1182
6.4705 | -34.3625 | -12.1281
6.4766 | -34.3629 | -12.1385
6.4827 | -34.3632 | -12.1489
6.4888 | -34.3636 | -12.1591
6.4949 | -34.364 | -12.1691
6.501 | -34.3644 | -12.1795
6.5071 | -34.3647 | -12.1896
6.5132 | -34.3651 | -12.2
6.5194 | -34.3655 | -12.2103
6.5255 | -34.3659 | -12.2204
6.5316 | -34.3662 | -12.2309
6.5377 | -34.3666 | -12.2412
6.5438 | -34.3669 | -12.2512
6.5499 | -34.3673 | -12.2614
6.556 | -34.3677 | -12.2717
6.5621 | -34.368 | -12.2822
6.5682 | -34.3684 | -12.2926
6.5743 | -34.3689 | -12.3028
6.5804 | -34.3692 | -12.3132
6.5865 | -34.3696 | -12.3235
6.5926 | -34.37 | -12.3337
6.5987 | -34.3704 | -12.3437
6.6048 | -34.3707 | -12.354
6.6109 | -34.3711 | -12.3642
6.617 | -34.3714 | -12.3747
6.6231 | -34.3718 | -12.3849
6.6292 | -34.3722 | -12.3952
6.6353 | -34.3726 | -12.4058
6.6414 | -34.373 | -12.4159
6.6475 | -34.3734 | -12.4262
6.6536 | -34.3737 | -12.4363
6.6597 | -34.3741 | -12.4463
6.6659 | -34.3744 | -12.4566
6.672 | -34.3748 | -12.4667
6.6781 | -34.3752 | -12.477
6.6842 | -34.3756 | -12.487
6.6903 | -34.376 | -12.4973
6.6964 | -34.3764 | -12.5075
6.7025 | -34.3768 | -12.5176
6.7086 | -34.3771 | -12.5278
6.7147 | -34.3775 | -12.5381
6.7208 | -34.3778 | -12.5484
6.7269 | -34.3782 | -12.5586
6.733 | -34.3786 | -12.5686
6.7391 | -34.379 | -12.5788
6.7452 | -34.3793 | -12.5888
6.7513 | -34.3797 | -12.5987
6.7574 | -34.3801 | -12.609
6.7635 | -34.3805 | -12.619
6.7696 | -34.3809 | -12.629
6.7757 | -34.3813 | -12.6393
6.7818 | -34.3817 | -12.6495
6.7879 | -34.3821 | -12.6595
6.794 | -34.3824 | -12.6694
6.8001 | -34.3828 | -12.6795
6.8063 | -34.3831 | -12.6897
6.8124 | -34.3835 | -12.7
6.8185 | -34.3838 | -12.7104
6.8246 | -34.3842 | -12.7203
6.8307 | -34.3846 | -12.7304
6.8368 | -34.385 | -12.7408
6.8429 | -34.3854 | -12.7509
6.849 | -34.3857 | -12.7611
6.8551 | -34.3862 | -12.7712
6.8612 | -34.3865 | -12.7813
6.8673 | -34.3869 | -12.7914
6.8734 | -34.3873 | -12.8016
6.8795 | -34.3876 | -12.8118
6.8856 | -34.388 | -12.822
6.8917 | -34.3884 | -12.832
6.8978 | -34.3887 | -12.8419
6.9039 | -34.3891 | -12.8517
6.91 | -34.3895 | -12.8619
6.9161 | -34.3899 | -12.8718
6.9222 | -34.3903 | -12.8819
6.9283 | -34.3907 | -12.8917
6.9344 | -34.3911 | -12.9019
6.9405 | -34.3915 | -12.9119
6.9466 | -34.3919 | -12.9223
6.9528 | -34.3923 | -12.9325
6.9589 | -34.3927 | -12.9428
6.965 | -34.3931 | -12.953
6.9711 | -34.3935 | -12.9632
6.9772 | -34.3939 | -12.9733
6.9833 | -34.3943 | -12.9834
6.9894 | -34.3947 | -12.9938
6.9955 | -34.3951 | -13.0036
7.0016 | -34.3954 | -13.0137
7.0077 | -34.3958 | -13.0237
7.0138 | -34.3962 | -13.0338
7.0199 | -34.3966 | -13.0437
7.026 | -34.397 | -13.0538
7.0321 | -34.3974 | -13.0636
7.0382 | -34.3978 | -13.0734
7.0443 | -34.3982 | -13.0836
7.0504 | -34.3986 | -13.0938
7.0565 | -34.399 | -13.1039
7.0626 | -34.3994 | -13.1137
7.0687 | -34.3998 | -13.1234
7.0748 | -34.4002 | -13.1331
7.0809 | -34.4006 | -13.1432
7.087 | -34.401 | -13.1533
7.0932 | -34.4013 | -13.1634
7.0993 | -34.4017 | -13.1734
7.1054 | -34.4021 | -13.1832
7.1115 | -34.4025 | -13.193
7.1176 | -34.4029 | -13.203
7.1237 | -34.4033 | -13.213
7.1298 | -34.4037 | -13.2229
7.1359 | -34.404 | -13.2329
7.142 | -34.4044 | -13.2426
7.1481 | -34.4048 | -13.2526
7.1542 | -34.4052 | -13.2626
7.1603 | -34.4056 | -13.2726
7.1664 | -34.406 | -13.2825
7.1725 | -34.4064 | -13.2926
7.1786 | -34.4069 | -13.3026
7.1847 | -34.4072 | -13.3126
7.1908 | -34.4076 | -13.3228
7.1969 | -34.408 | -13.333
7.203 | -34.4084 | -13.3428
7.2091 | -34.4088 | -13.3527
7.2152 | -34.4092 | -13.3627
7.2213 | -34.4096 | -13.3725
7.2274 | -34.41 | -13.3824
7.2335 | -34.4104 | -13.3924
7.2397 | -34.4108 | -13.4023
7.2458 | -34.4112 | -13.412
7.2519 | -34.4116 | -13.422
7.258 | -34.412 | -13.4322
7.2641 | -34.4124 | -13.4422
7.2702 | -34.4128 | -13.4523
7.2763 | -34.4132 | -13.4623
7.2824 | -34.4136 | -13.4725
7.2885 | -34.414 | -13.4825
7.2946 | -34.4144 | -13.4925
7.3007 | -34.4148 | -13.5026
7.3068 | -34.4152 | -13.5128
7.3129 | -34.4156 | -13.523
7.319 | -34.416 | -13.533
7.3251 | -34.4164 | -13.5427
7.3312 | -34.4168 | -13.5527
7.3373 | -34.4172 | -13.5625
7.3434 | -34.4176 | -13.5723
7.3495 | -34.4181 | -13.5822
7.3556 | -34.4185 | -13.5923
7.3617 | -34.4189 | -13.6024
7.3678 | -34.4193 | -13.6123
7.3739 | -34.4196 | -13.6225
7.3801 | -34.42 | -13.6327
7.3862 | -34.4205 | -13.6427
7.3923 | -34.4209 | -13.6528
7.3984 | -34.4213 | -13.6625
7.4045 | -34.4217 | -13.6725
7.4106 | -34.4221 | -13.6824
7.4167 | -34.4225 | -13.6924
7.4228 | -34.423 | -13.7026
7.4289 | -34.4234 | -13.7125
7.435 | -34.4238 | -13.7222
7.4411 | -34.4242 | -13.7323
7.4472 | -34.4246 | -13.7424
7.4533 | -34.425 | -13.7521
7.4594 | -34.4254 | -13.7624
7.4655 | -34.4258 | -13.7722
7.4716 | -34.4262 | -13.7823
7.4777 | -34.4266 | -13.7923
7.4838 | -34.427 | -13.8021
7.4899 | -34.4274 | -13.812
7.496 | -34.4278 | -13.8218
7.5021 | -34.4283 | -13.8318
7.5082 | -34.4287 | -13.8417
7.5143 | -34.4291 | -13.8518
7.5204 | -34.4295 | -13.8619
7.5266 | -34.4299 | -13.8717
7.5327 | -34.4303 | -13.8816
7.5388 | -34.4307 | -13.8919
7.5449 | -34.4311 | -13.9021
7.551 | -34.4315 | -13.9119
7.5571 | -34.432 | -13.922
7.5632 | -34.4324 | -13.9321
7.5693 | -34.4328 | -13.9419
7.5754 | -34.4332 | -13.9518
7.5815 | -34.4336 | -13.9618
7.5876 | -34.434 | -13.972
7.5937 | -34.4345 | -13.9822
7.5998 | -34.4349 | -13.9921
7.6059 | -34.4354 | -14.0023
7.612 | -34.4358 | -14.0122
7.6181 | -34.4362 | -14.0221
7.6242 | -34.4366 | -14.0324
7.6303 | -34.437 | -14.0426
7.6364 | -34.4374 | -14.0527
7.6425 | -34.4378 | -14.0628
7.6486 | -34.4382 | -14.0728
7.6547 | -34.4387 | -14.0827
7.6608 | -34.4391 | -14.0926
7.667 | -34.4395 | -14.1028
7.6731 | -34.4399 | -14.1129
7.6792 | -34.4403 | -14.1231
7.6853 | -34.4407 | -14.1332
7.6914 | -34.4412 | -14.1433
7.6975 | -34.4416 | -14.1535
7.7036 | -34.442 | -14.1636
7.7097 | -34.4424 | -14.1737
7.7158 | -34.4429 | -14.1839
7.7219 | -34.4433 | -14.1938
7.728 | -34.4437 | -14.2039
7.7341 | -34.4441 | -14.2141
7.7402 | -34.4446 | -14.2242
7.7463 | -34.445 | -14.2343
7.7524 | -34.4454 | -14.2446
7.7585 | -34.4458 | -14.2547
7.7646 | -34.4462 | -14.2647
7.7707 | -34.4466 | -14.275
7.7768 | -34.447 | -14.2852
7.7829 | -34.4474 | -14.2956
7.789 | -34.4479 | -14.3059
7.7951 | -34.4483 | -14.3159
7.8012 | -34.4487 | -14.326
7.8073 | -34.4492 | -14.3364
7.8135 | -34.4496 | -14.3464
7.8196 | -34.45 | -14.3568
7.8257 | -34.4504 | -14.3671
7.8318 | -34.4508 | -14.3773
7.8379 | -34.4513 | -14.3873
7.844 | -34.4517 | -14.3973
7.8501 | -34.4521 | -14.4075
7.8562 | -34.4525 | -14.4179
7.8623 | -34.453 | -14.4279
7.8684 | -34.4534 | -14.4381
7.8745 | -34.4538 | -14.4483
7.8806 | -34.4542 | -14.4586
7.8867 | -34.4547 | -14.4684
7.8928 | -34.4551 | -14.4783
7.8989 | -34.4555 | -14.4885
7.905 | -34.4559 | -14.4987
7.9111 | -34.4563 | -14.5089
7.9172 | -34.4568 | -14.519
7.9233 | -34.4571 | -14.5292
7.9294 | -34.4576 | -14.5394
7.9355 | -34.458 | -14.5493
7.9416 | -34.4584 | -14.5595
7.9477 | -34.4588 | -14.5695
7.9539 | -34.4592 | -14.5795
7.96 | -34.4597 | -14.5898
7.9661 | -34.4601 | -14.6
7.9722 | -34.4606 | -14.61
7.9783 | -34.461 | -14.6204
7.9844 | -34.4614 | -14.6305
7.9905 | -34.4618 | -14.6408
7.9966 | -34.4623 | -14.6511
8.0027 | -34.4627 | -14.6612
8.0088 | -34.4631 | -14.6713
8.0149 | -34.4635 | -14.6815
8.021 | -34.4639 | -14.6917
8.0271 | -34.4643 | -14.7022
8.0332 | -34.4647 | -14.7123
8.0393 | -34.4652 | -14.7226
8.0454 | -34.4656 | -14.733
8.0515 | -34.4661 | -14.7431
8.0576 | -34.4665 | -14.7536
8.0637 | -34.4669 | -14.7637
8.0698 | -34.4673 | -14.7737
8.0759 | -34.4677 | -14.7837
8.082 | -34.4681 | -14.7937
8.0881 | -34.4685 | -14.804
8.0942 | -34.469 | -14.8142
8.1004 | -34.4695 | -14.8245
8.1065 | -34.4699 | -14.8345
8.1126 | -34.4702 | -14.8447
8.1187 | -34.4707 | -14.855
8.1248 | -34.4711 | -14.8649
8.1309 | -34.4715 | -14.8749
8.137 | -34.4719 | -14.8849
8.1431 | -34.4724 | -14.8953
8.1492 | -34.4728 | -14.9054
8.1553 | -34.4732 | -14.9156
8.1614 | -34.4736 | -14.926
8.1675 | -34.474 | -14.9361
8.1736 | -34.4744 | -14.9464
8.1797 | -34.4748 | -14.9567
8.1858 | -34.4753 | -14.9669
8.1919 | -34.4757 | -14.9771
8.198 | -34.4761 | -14.987
8.2041 | -34.4766 | -14.9972
8.2102 | -34.477 | -15.0074
8.2163 | -34.4774 | -15.0174
8.2224 | -34.4778 | -15.0275
8.2285 | -34.4783 | -15.0374
8.2346 | -34.4787 | -15.0475
8.2408 | -34.4791 | -15.0575
8.2469 | -34.4795 | -15.0677
8.253 | -34.48 | -15.0777
8.2591 | -34.4804 | -15.088
8.2652 | -34.4809 | -15.0982
8.2713 | -34.4813 | -15.1084
8.2774 | -34.4817 | -15.1186
8.2835 | -34.4821 | -15.1291
8.2896 | -34.4826 | -15.139
8.2957 | -34.483 | -15.1492
8.3018 | -34.4834 | -15.1594
8.3079 | -34.4839 | -15.1694
8.314 | -34.4843 | -15.1795
8.3201 | -34.4847 | -15.1896
8.3262 | -34.4851 | -15.1998
8.3323 | -34.4856 | -15.2099
8.3384 | -34.486 | -15.2202
8.3445 | -34.4864 | -15.2302
8.3506 | -34.4869 | -15.2404
8.3567 | -34.4873 | -15.2507
8.3628 | -34.4877 | -15.2608
8.3689 | -34.4881 | -15.2708
8.375 | -34.4886 | -15.281
8.3812 | -34.489 | -15.291
8.3873 | -34.4894 | -15.3012
8.3934 | -34.4898 | -15.3114
8.3995 | -34.4902 | -15.3214
8.4056 | -34.4907 | -15.3315
8.4117 | -34.4911 | -15.3417
8.4178 | -34.4915 | -15.3519
8.4239 | -34.4919 | -15.3619
8.43 | -34.4923 | -15.372
8.4361 | -34.4927 | -15.3822
8.4422 | -34.4932 | -15.3922
8.4483 | -34.4936 | -15.4022
8.4544 | -34.494 | -15.4124
8.4605 | -34.4945 | -15.4224
8.4666 | -34.4949 | -15.4327
8.4727 | -34.4953 | -15.4429
8.4788 | -34.4957 | -15.453
8.4849 | -34.4961 | -15.4631
8.491 | -34.4965 | -15.4732
8.4971 | -34.497 | -15.4831
8.5032 | -34.4974 | -15.493
8.5093 | -34.4978 | -15.503
8.5154 | -34.4982 | -15.5132
8.5215 | -34.4987 | -15.5231
8.5277 | -34.4991 | -15.5332
8.5338 | -34.4995 | -15.5432
8.5399 | -34.4999 | -15.5533
8.546 | -34.5004 | -15.5632
8.5521 | -34.5008 | -15.5733
8.5582 | -34.5012 | -15.5833
8.5643 | -34.5016 | -15.5931
8.5704 | -34.5021 | -15.6029
8.5765 | -34.5025 | -15.6131
8.5826 | -34.5029 | -15.6228
8.5887 | -34.5034 | -15.6328
8.5948 | -34.5038 | -15.6429
8.6009 | -34.5042 | -15.6528
8.607 | -34.5047 | -15.663
8.6131 | -34.5051 | -15.6731
8.6192 | -34.5055 | -15.6829
8.6253 | -34.506 | -15.6928
8.6314 | -34.5064 | -15.7028
8.6375 | -34.5068 | -15.7129
8.6436 | -34.5072 | -15.7229
8.6497 | -34.5076 | -15.7328
8.6558 | -34.508 | -15.7427
8.6619 | -34.5084 | -15.7528
8.6681 | -34.5088 | -15.7629
8.6742 | -34.5093 | -15.7729
8.6803 | -34.5097 | -15.7829
8.6864 | -34.5101 | -15.7926
8.6925 | -34.5106 | -15.8024
8.6986 | -34.511 | -15.8124
8.7047 | -34.5114 | -15.8226
8.7108 | -34.5118 | -15.8326
8.7169 | -34.5122 | -15.8426
8.723 | -34.5126 | -15.8524
8.7291 | -34.5131 | -15.8622
8.7352 | -34.5135 | -15.8724
8.7413 | -34.5139 | -15.8824
8.7474 | -34.5143 | -15.8922
8.7535 | -34.5148 | -15.9022
8.7596 | -34.5152 | -15.9121
8.7657 | -34.5156 | -15.9222
8.7718 | -34.516 | -15.9322
8.7779 | -34.5165 | -15.9423
8.784 | -34.5169 | -15.9523
8.7901 | -34.5174 | -15.9622
8.7962 | -34.5178 | -15.9719
8.8023 | -34.5182 | -15.9818
8.8084 | -34.5186 | -15.9917
8.8146 | -34.5191 | -16.0016
8.8207 | -34.5194 | -16.0115
8.8268 | -34.5199 | -16.0211
8.8329 | -34.5203 | -16.031
8.839 | -34.5208 | -16.041
8.8451 | -34.5212 | -16.0508
8.8512 | -34.5216 | -16.0606
8.8573 | -34.522 | -16.0706
8.8634 | -34.5225 | -16.0802
8.8695 | -34.5229 | -16.0903
8.8756 | -34.5233 | -16.1002
8.8817 | -34.5237 | -16.11
8.8878 | -34.5241 | -16.1201
8.8939 | -34.5246 | -16.13
8.9 | -34.525 | -16.1398
8.9061 | -34.5254 | -16.1497
8.9122 | -34.5259 | -16.1596
8.9183 | -34.5263 | -16.1695
8.9244 | -34.5268 | -16.1795
8.9305 | -34.5272 | -16.1892
8.9366 | -34.5276 | -16.1991
8.9427 | -34.5281 | -16.209
8.9488 | -34.5285 | -16.2188
8.955 | -34.5289 | -16.2287
8.9611 | -34.5294 | -16.2387
8.9672 | -34.5299 | -16.2485
8.9733 | -34.5303 | -16.2582
8.9794 | -34.5307 | -16.2679
8.9855 | -34.5312 | -16.2778
8.9916 | -34.5316 | -16.2879
8.9977 | -34.532 | -16.2977
9.0038 | -34.5325 | -16.3074
9.0099 | -34.533 | -16.3173
9.016 | -34.5334 | -16.3273
9.0221 | -34.5338 | -16.3369
9.0282 | -34.5343 | -16.3466
9.0343 | -34.5347 | -16.3565
9.0404 | -34.5352 | -16.3664
9.0465 | -34.5356 | -16.3764
9.0526 | -34.5361 | -16.3861
9.0587 | -34.5366 | -16.396
9.0648 | -34.537 | -16.4057
9.0709 | -34.5374 | -16.4155
9.077 | -34.5379 | -16.4254
9.0831 | -34.5383 | -16.4354
9.0892 | -34.5388 | -16.4453
9.0953 | -34.5392 | -16.4552
9.1015 | -34.5397 | -16.465
9.1076 | -34.5401 | -16.4747
9.1137 | -34.5406 | -16.4845
9.1198 | -34.541 | -16.4943
9.1259 | -34.5415 | -16.5041
9.132 | -34.5419 | -16.5141
9.1381 | -34.5424 | -16.5238
9.1442 | -34.5429 | -16.5333
9.1503 | -34.5433 | -16.5432
9.1564 | -34.5438 | -16.553
9.1625 | -34.5442 | -16.5628
9.1686 | -34.5447 | -16.5725
9.1747 | -34.5451 | -16.5823
9.1808 | -34.5456 | -16.5921
9.1869 | -34.5461 | -16.6019
9.193 | -34.5465 | -16.6117
9.1991 | -34.5469 | -16.6215
9.2052 | -34.5474 | -16.6313
9.2113 | -34.5479 | -16.6411
9.2174 | -34.5484 | -16.6508
9.2235 | -34.5488 | -16.6606
9.2296 | -34.5493 | -16.6704
9.2357 | -34.5497 | -16.6803
9.2419 | -34.5502 | -16.6901
9.248 | -34.5507 | -16.6998
9.2541 | -34.5512 | -16.7096
9.2602 | -34.5516 | -16.7196
9.2663 | -34.5521 | -16.729
9.2724 | -34.5526 | -16.7388
9.2785 | -34.553 | -16.7485
9.2846 | -34.5535 | -16.7581
9.2907 | -34.5539 | -16.7678
9.2968 | -34.5544 | -16.7777
9.3029 | -34.5548 | -16.7876
9.309 | -34.5553 | -16.7974
9.3151 | -34.5557 | -16.8071
9.3212 | -34.5563 | -16.8167
9.3273 | -34.5567 | -16.8264
9.3334 | -34.5572 | -16.836
9.3395 | -34.5576 | -16.8456
9.3456 | -34.5581 | -16.8552
9.3517 | -34.5585 | -16.865
9.3578 | -34.559 | -16.8747
9.3639 | -34.5595 | -16.8843
9.37 | -34.5599 | -16.8941
9.3761 | -34.5604 | -16.904
9.3822 | -34.5609 | -16.9137
9.3884 | -34.5613 | -16.9232
9.3945 | -34.5618 | -16.933
9.4006 | -34.5623 | -16.9427
9.4067 | -34.5628 | -16.9525
9.4128 | -34.5632 | -16.9624
9.4189 | -34.5637 | -16.9723
9.425 | -34.5642 | -16.9819
9.4311 | -34.5647 | -16.9916
9.4372 | -34.5652 | -17.0013
9.4433 | -34.5656 | -17.011
9.4494 | -34.5661 | -17.0206
9.4555 | -34.5666 | -17.0304
9.4616 | -34.5671 | -17.0401
9.4677 | -34.5675 | -17.0498
9.4738 | -34.568 | -17.0596
9.4799 | -34.5685 | -17.0693
9.486 | -34.5689 | -17.079
9.4921 | -34.5694 | -17.0887
9.4982 | -34.5699 | -17.0982
9.5043 | -34.5704 | -17.1079
9.5104 | -34.5708 | -17.1174
9.5165 | -34.5713 | -17.1271
9.5226 | -34.5718 | -17.1369
9.5288 | -34.5722 | -17.1464
9.5349 | -34.5727 | -17.1561
9.541 | -34.5732 | -17.1656
9.5471 | -34.5736 | -17.1753
9.5532 | -34.5741 | -17.185
9.5593 | -34.5746 | -17.1948
9.5654 | -34.5751 | -17.2044
9.5715 | -34.5755 | -17.214
9.5776 | -34.576 | -17.2239
9.5837 | -34.5765 | -17.2336
9.5898 | -34.5769 | -17.2434
9.5959 | -34.5774 | -17.2531
9.602 | -34.5778 | -17.2626
9.6081 | -34.5783 | -17.2723
9.6142 | -34.5788 | -17.2819
9.6203 | -34.5793 | -17.2916
9.6264 | -34.5797 | -17.3011
9.6325 | -34.5802 | -17.3109
9.6386 | -34.5807 | -17.3204
9.6447 | -34.5811 | -17.3301
9.6508 | -34.5816 | -17.3396
9.6569 | -34.5821 | -17.3493
9.663 | -34.5826 | -17.3589
9.6691 | -34.5831 | -17.3686
9.6753 | -34.5835 | -17.3783
9.6814 | -34.584 | -17.3878
9.6875 | -34.5845 | -17.3972
9.6936 | -34.5849 | -17.4067
9.6997 | -34.5854 | -17.4165
9.7058 | -34.586 | -17.426
9.7119 | -34.5864 | -17.4358
9.718 | -34.5869 | -17.4456
9.7241 | -34.5874 | -17.4551
9.7302 | -34.5878 | -17.4648
9.7363 | -34.5883 | -17.4744
9.7424 | -34.5888 | -17.4839
9.7485 | -34.5893 | -17.4935
9.7546 | -34.5897 | -17.503
9.7607 | -34.5902 | -17.5129
9.7668 | -34.5907 | -17.5226
9.7729 | -34.5912 | -17.5323
9.779 | -34.5917 | -17.5417
9.7851 | -34.5922 | -17.5514
9.7912 | -34.5927 | -17.561
9.7973 | -34.5932 | -17.5708
9.8034 | -34.5937 | -17.5802
9.8095 | -34.5941 | -17.59
9.8157 | -34.5946 | -17.5997
9.8218 | -34.5951 | -17.6095
9.8279 | -34.5956 | -17.6189
9.834 | -34.596 | -17.6287
9.8401 | -34.5965 | -17.6383
9.8462 | -34.597 | -17.6476
9.8523 | -34.5975 | -17.6572
9.8584 | -34.5979 | -17.6669
9.8645 | -34.5984 | -17.6766
9.8706 | -34.5989 | -17.6862
9.8767 | -34.5993 | -17.6957
9.8828 | -34.5998 | -17.7052
9.8889 | -34.6002 | -17.7146
9.895 | -34.6007 | -17.7242
9.9011 | -34.6012 | -17.7338
9.9072 | -34.6017 | -17.7434
9.9133 | -34.6021 | -17.753
9.9194 | -34.6026 | -17.7625
9.9255 | -34.6031 | -17.772
9.9316 | -34.6036 | -17.7816
9.9377 | -34.604 | -17.7911
9.9438 | -34.6045 | -17.8007
9.9499 | -34.605 | -17.8106
9.956 | -34.6054 | -17.8202
9.9622 | -34.6059 | -17.8298
9.9683 | -34.6064 | -17.8394
9.9744 | -34.6068 | -17.8491
9.9805 | -34.6074 | -17.8586
9.9866 | -34.6078 | -17.8682
9.9927 | -34.6083 | -17.8777
9.9988 | -34.6087 | -17.8873
10.0049 | -34.6092 | -17.897
10.011 | -34.6097 | -17.9066
10.0171 | -34.6102 | -17.9161
10.0232 | -34.6107 | -17.9255
10.0293 | -34.6112 | -17.9351
10.0354 | -34.6117 | -17.9446
10.0415 | -34.6122 | -17.9543
10.0476 | -34.6126 | -17.9639
10.0537 | -34.6131 | -17.9736
10.0598 | -34.6136 | -17.9832
10.0659 | -34.614 | -17.9928
10.072 | -34.6145 | -18.0022
10.0781 | -34.615 | -18.0117
10.0842 | -34.6155 | -18.0213
10.0903 | -34.6159 | -18.0309
10.0964 | -34.6164 | -18.0402
10.1026 | -34.6169 | -18.0498
10.1087 | -34.6174 | -18.0594
10.1148 | -34.6178 | -18.0691
10.1209 | -34.6183 | -18.0786
10.127 | -34.6188 | -18.0879
10.1331 | -34.6193 | -18.0974
10.1392 | -34.6197 | -18.1067
10.1453 | -34.6202 | -18.1162
10.1514 | -34.6207 | -18.1258
10.1575 | -34.6211 | -18.1356
10.1636 | -34.6216 | -18.1449
10.1697 | -34.6221 | -18.1545
10.1758 | -34.6226 | -18.1639
10.1819 | -34.6231 | -18.1735
10.188 | -34.6235 | -18.183
10.1941 | -34.624 | -18.1926
10.2002 | -34.6244 | -18.2022
10.2063 | -34.6249 | -18.2116
10.2124 | -34.6254 | -18.2213
10.2185 | -34.6259 | -18.2308
10.2246 | -34.6264 | -18.2404
10.2307 | -34.6268 | -18.2499
10.2368 | -34.6273 | -18.2592
10.2429 | -34.6278 | -18.2687
10.2491 | -34.6282 | -18.2783
10.2552 | -34.6287 | -18.2879
10.2613 | -34.6292 | -18.2974
10.2674 | -34.6296 | -18.3069
10.2735 | -34.6301 | -18.3165
10.2796 | -34.6306 | -18.326
10.2857 | -34.6311 | -18.3357
10.2918 | -34.6315 | -18.3454
10.2979 | -34.632 | -18.355
10.304 | -34.6325 | -18.3643
10.3101 | -34.633 | -18.3738
10.3162 | -34.6334 | -18.3835
10.3223 | -34.6339 | -18.3928
10.3284 | -34.6344 | -18.4022
10.3345 | -34.6348 | -18.4117
10.3406 | -34.6353 | -18.4211
10.3467 | -34.6358 | -18.4306
10.3528 | -34.6362 | -18.4399
10.3589 | -34.6367 | -18.4495
10.365 | -34.6372 | -18.4591
10.3711 | -34.6377 | -18.4686
10.3772 | -34.6381 | -18.478
10.3833 | -34.6386 | -18.4876
10.3895 | -34.6391 | -18.4971
10.3956 | -34.6396 | -18.5065
10.4017 | -34.64 | -18.5159
10.4078 | -34.6405 | -18.5254
10.4139 | -34.6409 | -18.5349
10.42 | -34.6414 | -18.5447
10.4261 | -34.6419 | -18.5542
10.4322 | -34.6424 | -18.5637
10.4383 | -34.6428 | -18.5731
10.4444 | -34.6433 | -18.5824
10.4505 | -34.6437 | -18.5919
10.4566 | -34.6442 | -18.6013
10.4627 | -34.6446 | -18.6107
10.4688 | -34.6451 | -18.62
10.4749 | -34.6456 | -18.6295
10.481 | -34.6461 | -18.639
10.4871 | -34.6466 | -18.6483
10.4932 | -34.647 | -18.6577
10.4993 | -34.6475 | -18.6672
10.5054 | -34.648 | -18.6766
10.5115 | -34.6484 | -18.6861
10.5176 | -34.6489 | -18.6955
10.5237 | -34.6494 | -18.7051
10.5298 | -34.6498 | -18.7144
10.536 | -34.6503 | -18.7239
10.5421 | -34.6508 | -18.7333
10.5482 | -34.6512 | -18.7428
10.5543 | -34.6517 | -18.7523
10.5604 | -34.6522 | -18.7615
10.5665 | -34.6527 | -18.7708
10.5726 | -34.6531 | -18.7803
10.5787 | -34.6536 | -18.7896
10.5848 | -34.6541 | -18.7989
10.5909 | -34.6546 | -18.8086
10.597 | -34.655 | -18.8181
10.6031 | -34.6555 | -18.8278
10.6092 | -34.6559 | -18.8372
10.6153 | -34.6564 | -18.8467
10.6214 | -34.6569 | -18.856
10.6275 | -34.6574 | -18.8655
10.6336 | -34.6579 | -18.875
10.6397 | -34.6583 | -18.8845
10.6458 | -34.6588 | -18.894
10.6519 | -34.6593 | -18.9035
10.658 | -34.6597 | -18.913
10.6641 | -34.6602 | -18.9225
10.6702 | -34.6606 | -18.932
10.6764 | -34.6611 | -18.9416
10.6825 | -34.6616 | -18.9508
10.6886 | -34.6621 | -18.9602
10.6947 | -34.6625 | -18.9696
10.7008 | -34.663 | -18.9791
10.7069 | -34.6635 | -18.9886
10.713 | -34.664 | -18.998
10.7191 | -34.6644 | -19.0074
10.7252 | -34.6649 | -19.0165
10.7313 | -34.6653 | -19.0259
10.7374 | -34.6658 | -19.0353
10.7435 | -34.6663 | -19.0448
10.7496 | -34.6667 | -19.054
10.7557 | -34.6672 | -19.0636
10.7618 | -34.6677 | -19.0727
10.7679 | -34.6681 | -19.0822
10.774 | -34.6686 | -19.0916
10.7801 | -34.6691 | -19.1011
10.7862 | -34.6695 | -19.1103
10.7923 | -34.67 | -19.1199
10.7984 | -34.6705 | -19.1291
10.8045 | -34.6709 | -19.1383
10.8106 | -34.6714 | -19.1478
10.8168 | -34.6719 | -19.1571
10.8229 | -34.6723 | -19.1665
10.829 | -34.6728 | -19.1758
10.8351 | -34.6733 | -19.1849
10.8412 | -34.6738 | -19.1942
10.8473 | -34.6743 | -19.2034
10.8534 | -34.6747 | -19.2128
10.8595 | -34.6752 | -19.2221
10.8656 | -34.6757 | -19.2313
10.8717 | -34.6761 | -19.2407
10.8778 | -34.6766 | -19.2501
10.8839 | -34.6771 | -19.2595
10.89 | -34.6776 | -19.2689
10.8961 | -34.6781 | -19.2782
10.9022 | -34.6785 | -19.2875
10.9083 | -34.679 | -19.2968
10.9144 | -34.6794 | -19.306
10.9205 | -34.6799 | -19.3154
10.9266 | -34.6804 | -19.3247
10.9327 | -34.6809 | -19.334
10.9388 | -34.6814 | -19.3432
10.9449 | -34.6818 | -19.3524
10.951 | -34.6823 | -19.3617
10.9571 | -34.6828 | -19.3711
10.9633 | -34.6832 | -19.3805
10.9694 | -34.6837 | -19.3901
10.9755 | -34.6842 | -19.3995
10.9816 | -34.6846 | -19.4089
10.9877 | -34.6851 | -19.4181
10.9938 | -34.6856 | -19.4275
10.9999 | -34.6861 | -19.4367
11.006 | -34.6865 | -19.4459
11.0121 | -34.687 | -19.4553
11.0182 | -34.6875 | -19.4646
11.0243 | -34.6879 | -19.4739
11.0304 | -34.6884 | -19.4832
11.0365 | -34.6889 | -19.4924
11.0426 | -34.6894 | -19.5017
11.0487 | -34.6899 | -19.5111
11.0548 | -34.6903 | -19.5203
11.0609 | -34.6908 | -19.5297
11.067 | -34.6913 | -19.5391
11.0731 | -34.6918 | -19.5483
11.0792 | -34.6923 | -19.5575
11.0853 | -34.6927 | -19.5668
11.0914 | -34.6932 | -19.5762
11.0975 | -34.6937 | -19.5855
11.1037 | -34.6942 | -19.5949
11.1098 | -34.6946 | -19.6042
11.1159 | -34.6951 | -19.6134
11.122 | -34.6956 | -19.6226
11.1281 | -34.6961 | -19.6319
11.1342 | -34.6965 | -19.6412
11.1403 | -34.697 | -19.6506
11.1464 | -34.6975 | -19.6601
11.1525 | -34.698 | -19.6692
11.1586 | -34.6985 | -19.6785
11.1647 | -34.6989 | -19.6877
11.1708 | -34.6994 | -19.697
11.1769 | -34.6999 | -19.7063
11.183 | -34.7003 | -19.7154
11.1891 | -34.7008 | -19.7247
11.1952 | -34.7013 | -19.734
11.2013 | -34.7018 | -19.7433
11.2074 | -34.7023 | -19.7526
11.2135 | -34.7027 | -19.7621
11.2196 | -34.7032 | -19.7714
11.2257 | -34.7037 | -19.7808
11.2318 | -34.7042 | -19.7899
11.2379 | -34.7047 | -19.7993
11.244 | -34.7052 | -19.8088
11.2502 | -34.7057 | -19.818
11.2563 | -34.7062 | -19.8274
11.2624 | -34.7067 | -19.8366
11.2685 | -34.7071 | -19.8458
11.2746 | -34.7076 | -19.855
11.2807 | -34.7081 | -19.8645
11.2868 | -34.7086 | -19.8739
11.2929 | -34.7091 | -19.8828
11.299 | -34.7096 | -19.892
11.3051 | -34.71 | -19.9015
11.3112 | -34.7105 | -19.9109
11.3173 | -34.711 | -19.92
11.3234 | -34.7115 | -19.9291
11.3295 | -34.712 | -19.9385
11.3356 | -34.7125 | -19.9479
11.3417 | -34.7129 | -19.957
11.3478 | -34.7134 | -19.9665
11.3539 | -34.7139 | -19.9759
11.36 | -34.7143 | -19.985
11.3661 | -34.7148 | -19.9941
11.3722 | -34.7153 | -20.0032
11.3783 | -34.7158 | -20.0126
11.3844 | -34.7163 | -20.0219
11.3906 | -34.7168 | -20.0312
11.3967 | -34.7173 | -20.0404
11.4028 | -34.7178 | -20.0498
11.4089 | -34.7183 | -20.0589
11.415 | -34.7187 | -20.0681
11.4211 | -34.7192 | -20.0773
11.4272 | -34.7197 | -20.0866
11.4333 | -34.7201 | -20.0957
11.4394 | -34.7206 | -20.1049
11.4455 | -34.7211 | -20.114
11.4516 | -34.7216 | -20.1233
11.4577 | -34.7221 | -20.1325
11.4638 | -34.7226 | -20.1416
11.4699 | -34.7231 | -20.1509
11.476 | -34.7235 | -20.1605
11.4821 | -34.724 | -20.1697
11.4882 | -34.7245 | -20.179
11.4943 | -34.725 | -20.1883
11.5004 | -34.7255 | -20.1973
11.5065 | -34.726 | -20.2065
11.5126 | -34.7265 | -20.2156
11.5187 | -34.727 | -20.2249
11.5248 | -34.7275 | -20.2343
11.5309 | -34.728 | -20.2435
11.5371 | -34.7285 | -20.2526
11.5432 | -34.729 | -20.2619
11.5493 | -34.7294 | -20.2713
11.5554 | -34.7299 | -20.2805
11.5615 | -34.7304 | -20.2898
11.5676 | -34.7309 | -20.2991
11.5737 | -34.7314 | -20.3082
11.5798 | -34.7319 | -20.3172
11.5859 | -34.7324 | -20.3264
11.592 | -34.7329 | -20.3356
11.5981 | -34.7334 | -20.3447
11.6042 | -34.7339 | -20.3539
11.6103 | -34.7343 | -20.3629
11.6164 | -34.7348 | -20.372
11.6225 | -34.7353 | -20.3812
11.6286 | -34.7358 | -20.3904
11.6347 | -34.7363 | -20.3997
11.6408 | -34.7368 | -20.4087
11.6469 | -34.7373 | -20.4179
11.653 | -34.7378 | -20.427
11.6591 | -34.7382 | -20.4362
11.6652 | -34.7387 | -20.4453
11.6713 | -34.7392 | -20.4543
11.6775 | -34.7397 | -20.4634
11.6836 | -34.7402 | -20.4726
11.6897 | -34.7407 | -20.4818
11.6958 | -34.7412 | -20.4911
11.7019 | -34.7417 | -20.5005
11.708 | -34.7422 | -20.5096
11.7141 | -34.7427 | -20.5185
11.7202 | -34.7432 | -20.5277
11.7263 | -34.7437 | -20.5368
11.7324 | -34.7441 | -20.5462
11.7385 | -34.7446 | -20.5554
11.7446 | -34.7451 | -20.5646
11.7507 | -34.7456 | -20.5737
11.7568 | -34.7461 | -20.5829
11.7629 | -34.7466 | -20.5922
11.769 | -34.7471 | -20.6012
11.7751 | -34.7476 | -20.6103
11.7812 | -34.7481 | -20.6193
11.7873 | -34.7486 | -20.6285
11.7934 | -34.7491 | -20.6377
11.7995 | -34.7496 | -20.6468
11.8056 | -34.7501 | -20.6558
11.8117 | -34.7505 | -20.665
11.8178 | -34.7511 | -20.6739
11.824 | -34.7516 | -20.683
11.8301 | -34.752 | -20.6921
11.8362 | -34.7525 | -20.701
11.8423 | -34.753 | -20.7102
11.8484 | -34.7535 | -20.7194
11.8545 | -34.7541 | -20.7283
11.8606 | -34.7545 | -20.7373
11.8667 | -34.755 | -20.7466
11.8728 | -34.7555 | -20.7556
11.8789 | -34.756 | -20.7647
11.885 | -34.7565 | -20.7738
11.8911 | -34.757 | -20.7829
11.8972 | -34.7575 | -20.7919
11.9033 | -34.758 | -20.8009
11.9094 | -34.7585 | -20.8099
11.9155 | -34.759 | -20.8191
11.9216 | -34.7595 | -20.8281
11.9277 | -34.76 | -20.8371
11.9338 | -34.7605 | -20.8462
11.9399 | -34.761 | -20.8554
11.946 | -34.7615 | -20.8644
11.9521 | -34.7619 | -20.8734
11.9582 | -34.7625 | -20.8824
11.9644 | -34.7629 | -20.8913
11.9705 | -34.7635 | -20.9003
11.9766 | -34.764 | -20.9094
11.9827 | -34.7645 | -20.9184
11.9888 | -34.765 | -20.9273
11.9949 | -34.7656 | -20.9361
12.001 | -34.766 | -20.9453
12.0071 | -34.7665 | -20.9543
12.0132 | -34.767 | -20.9632
12.0193 | -34.7675 | -20.9724
12.0254 | -34.768 | -20.9814
12.0315 | -34.7685 | -20.9902
12.0376 | -34.769 | -20.9992
12.0437 | -34.7694 | -21.0082
12.0498 | -34.77 | -21.0173
12.0559 | -34.7705 | -21.0262
12.062 | -34.771 | -21.035
12.0681 | -34.7715 | -21.0441
12.0742 | -34.772 | -21.053
12.0803 | -34.7725 | -21.0622
12.0864 | -34.773 | -21.0713
12.0925 | -34.7735 | -21.0803
12.0986 | -34.774 | -21.0893
12.1047 | -34.7745 | -21.0983
12.1109 | -34.775 | -21.1072
12.117 | -34.7755 | -21.1161
12.1231 | -34.7761 | -21.1251
12.1292 | -34.7765 | -21.1341
12.1353 | -34.777 | -21.1431
12.1414 | -34.7776 | -21.1521
12.1475 | -34.778 | -21.1612
12.1536 | -34.7785 | -21.1702
12.1597 | -34.779 | -21.1792
12.1658 | -34.7795 | -21.1881
12.1719 | -34.78 | -21.1968
12.178 | -34.7805 | -21.2058
12.1841 | -34.781 | -21.2149
12.1902 | -34.7815 | -21.2239
12.1963 | -34.7821 | -21.2329
12.2024 | -34.7826 | -21.2416
12.2085 | -34.7831 | -21.2506
12.2146 | -34.7836 | -21.2597
12.2207 | -34.7841 | -21.2686
12.2268 | -34.7846 | -21.2776
12.2329 | -34.7851 | -21.2865
12.239 | -34.7857 | -21.2951
12.2451 | -34.7862 | -21.3041
12.2513 | -34.7867 | -21.3133
12.2574 | -34.7872 | -21.3222
12.2635 | -34.7877 | -21.3312
12.2696 | -34.7882 | -21.34
12.2757 | -34.7887 | -21.3487
12.2818 | -34.7893 | -21.3577
12.2879 | -34.7898 | -21.3667
12.294 | -34.7903 | -21.3754
12.3001 | -34.7908 | -21.3844
12.3062 | -34.7913 | -21.3933
12.3123 | -34.7918 | -21.402
12.3184 | -34.7923 | -21.4108
12.3245 | -34.7928 | -21.4196
12.3306 | -34.7933 | -21.4285
12.3367 | -34.7938 | -21.4374
12.3428 | -34.7943 | -21.4463
12.3489 | -34.7949 | -21.4549
12.355 | -34.7954 | -21.464
12.3611 | -34.7959 | -21.473
12.3672 | -34.7964 | -21.4819
12.3733 | -34.7969 | -21.4908
12.3794 | -34.7974 | -21.4993
12.3855 | -34.7979 | -21.5082
12.3916 | -34.7984 | -21.5171
12.3978 | -34.799 | -21.5261
12.4039 | -34.7995 | -21.535
12.41 | -34.8 | -21.5439
12.4161 | -34.8005 | -21.5524
12.4222 | -34.801 | -21.5613
12.4283 | -34.8015 | -21.5701
12.4344 | -34.8021 | -21.579
12.4405 | -34.8025 | -21.588
12.4466 | -34.803 | -21.5967
12.4527 | -34.8035 | -21.6055
12.4588 | -34.8041 | -21.6144
12.4649 | -34.8046 | -21.6233
12.471 | -34.8051 | -21.6323
12.4771 | -34.8056 | -21.6411
12.4832 | -34.8061 | -21.6499
12.4893 | -34.8066 | -21.6586
12.4954 | -34.8071 | -21.6673
12.5015 | -34.8076 | -21.6761
12.5076 | -34.8082 | -21.685
12.5137 | -34.8087 | -21.6939
12.5198 | -34.8092 | -21.7026
12.5259 | -34.8097 | -21.7114
12.532 | -34.8102 | -21.7202
12.5382 | -34.8107 | -21.7291
12.5443 | -34.8112 | -21.7379
12.5504 | -34.8117 | -21.7468
12.5565 | -34.8123 | -21.7556
12.5626 | -34.8128 | -21.7645
12.5687 | -34.8133 | -21.7734
12.5748 | -34.8138 | -21.7823
12.5809 | -34.8143 | -21.7911
12.587 | -34.8148 | -21.8
12.5931 | -34.8154 | -21.8087
12.5992 | -34.8159 | -21.8174
12.6053 | -34.8164 | -21.8263
12.6114 | -34.8169 | -21.835
12.6175 | -34.8174 | -21.8439
12.6236 | -34.8179 | -21.8528
12.6297 | -34.8185 | -21.8615
12.6358 | -34.819 | -21.8701
12.6419 | -34.8195 | -21.8789
12.648 | -34.82 | -21.8879
12.6541 | -34.8205 | -21.8966
12.6602 | -34.821 | -21.905
12.6663 | -34.8215 | -21.9138
12.6724 | -34.8221 | -21.9225
12.6785 | -34.8226 | -21.9312
12.6847 | -34.8231 | -21.9399
12.6908 | -34.8236 | -21.9486
12.6969 | -34.8241 | -21.9573
12.703 | -34.8246 | -21.9661
12.7091 | -34.8251 | -21.9749
12.7152 | -34.8256 | -21.9835
12.7213 | -34.8261 | -21.9924
12.7274 | -34.8267 | -22.0011
12.7335 | -34.8272 | -22.0096
12.7396 | -34.8277 | -22.0182
12.7457 | -34.8282 | -22.0272
12.7518 | -34.8287 | -22.0358
12.7579 | -34.8293 | -22.0447
12.764 | -34.8298 | -22.0536
12.7701 | -34.8303 | -22.0623
12.7762 | -34.8308 | -22.071
12.7823 | -34.8313 | -22.0797
12.7884 | -34.8319 | -22.0885
12.7945 | -34.8324 | -22.0974
12.8006 | -34.8329 | -22.106
12.8067 | -34.8334 | -22.1147
12.8128 | -34.8339 | -22.1233
12.8189 | -34.8345 | -22.132
12.8251 | -34.835 | -22.1408
12.8312 | -34.8355 | -22.1496
12.8373 | -34.836 | -22.1582
12.8434 | -34.8365 | -22.167
12.8495 | -34.837 | -22.1756
12.8556 | -34.8375 | -22.1842
12.8617 | -34.838 | -22.1929
12.8678 | -34.8385 | -22.2018
12.8739 | -34.8391 | -22.2104
12.88 | -34.8396 | -22.2191
12.8861 | -34.84 | -22.2276
12.8922 | -34.8406 | -22.2364
12.8983 | -34.8411 | -22.2452
12.9044 | -34.8416 | -22.2536
12.9105 | -34.8421 | -22.2622
12.9166 | -34.8426 | -22.2708
12.9227 | -34.8431 | -22.2795
12.9288 | -34.8436 | -22.2883
12.9349 | -34.8442 | -22.2969
12.941 | -34.8447 | -22.3056
12.9471 | -34.8452 | -22.3143
12.9532 | -34.8458 | -22.323
12.9593 | -34.8463 | -22.3319
12.9654 | -34.8468 | -22.3407
12.9716 | -34.8473 | -22.3495
12.9777 | -34.8479 | -22.3581
12.9838 | -34.8484 | -22.3668
12.9899 | -34.8489 | -22.3755
12.996 | -34.8494 | -22.3842
13.0021 | -34.85 | -22.3929
13.0082 | -34.8505 | -22.4014
13.0143 | -34.851 | -22.4098
13.0204 | -34.8515 | -22.4183
13.0265 | -34.8521 | -22.4269
13.0326 | -34.8526 | -22.4356
13.0387 | -34.8531 | -22.4444
13.0448 | -34.8536 | -22.4531
13.0509 | -34.8541 | -22.4617
13.057 | -34.8547 | -22.4703
13.0631 | -34.8552 | -22.4788
13.0692 | -34.8557 | -22.4874
13.0753 | -34.8563 | -22.4962
13.0814 | -34.8568 | -22.505
13.0875 | -34.8573 | -22.5134
13.0936 | -34.8578 | -22.5222
13.0997 | -34.8583 | -22.5307
13.1058 | -34.8589 | -22.5393
13.112 | -34.8594 | -22.548
13.1181 | -34.8599 | -22.5564
13.1242 | -34.8604 | -22.5651
13.1303 | -34.861 | -22.5738
13.1364 | -34.8615 | -22.5824
13.1425 | -34.862 | -22.591
13.1486 | -34.8625 | -22.5997
13.1547 | -34.863 | -22.6084
13.1608 | -34.8635 | -22.617
13.1669 | -34.8641 | -22.6254
13.173 | -34.8646 | -22.6341
13.1791 | -34.8651 | -22.6426
13.1852 | -34.8657 | -22.6512
13.1913 | -34.8662 | -22.6596
13.1974 | -34.8667 | -22.6683
13.2035 | -34.8672 | -22.6769
13.2096 | -34.8677 | -22.6854
13.2157 | -34.8683 | -22.694
13.2218 | -34.8688 | -22.7025
13.2279 | -34.8693 | -22.7108
13.234 | -34.8698 | -22.7194
13.2401 | -34.8704 | -22.728
13.2462 | -34.8709 | -22.7365
13.2524 | -34.8714 | -22.7452
13.2585 | -34.8719 | -22.7537
13.2646 | -34.8725 | -22.7622
13.2707 | -34.873 | -22.7709
13.2768 | -34.8735 | -22.7796
13.2829 | -34.8741 | -22.7883
13.289 | -34.8746 | -22.7968
13.2951 | -34.8751 | -22.8052
13.3012 | -34.8756 | -22.8138
13.3073 | -34.8762 | -22.8225
13.3134 | -34.8767 | -22.8311
13.3195 | -34.8772 | -22.8395
13.3256 | -34.8777 | -22.848
13.3317 | -34.8782 | -22.8565
13.3378 | -34.8787 | -22.8652
13.3439 | -34.8793 | -22.8739
13.35 | -34.8798 | -22.8826
13.3561 | -34.8803 | -22.891
13.3622 | -34.8808 | -22.8994
13.3683 | -34.8814 | -22.9078
13.3744 | -34.8819 | -22.9162
13.3805 | -34.8824 | -22.9249
13.3866 | -34.8829 | -22.9334
13.3927 | -34.8834 | -22.942
13.3989 | -34.8839 | -22.9505
13.405 | -34.8844 | -22.959
13.4111 | -34.885 | -22.9675
13.4172 | -34.8855 | -22.9761
13.4233 | -34.886 | -22.9848
13.4294 | -34.8865 | -22.9935
13.4355 | -34.887 | -23.0018
13.4416 | -34.8876 | -23.0102
13.4477 | -34.8881 | -23.0188
13.4538 | -34.8886 | -23.0274
13.4599 | -34.8892 | -23.036
13.466 | -34.8897 | -23.0446
13.4721 | -34.8902 | -23.053
13.4782 | -34.8907 | -23.0615
13.4843 | -34.8913 | -23.0701
13.4904 | -34.8918 | -23.0786
13.4965 | -34.8923 | -23.0872
13.5026 | -34.8928 | -23.0958
13.5087 | -34.8933 | -23.1042
13.5148 | -34.8939 | -23.1128
13.5209 | -34.8944 | -23.1215
13.527 | -34.8949 | -23.13
13.5331 | -34.8955 | -23.1385
13.5393 | -34.896 | -23.1468
13.5454 | -34.8965 | -23.1553
13.5515 | -34.897 | -23.164
13.5576 | -34.8976 | -23.1726
13.5637 | -34.8981 | -23.1812
13.5698 | -34.8986 | -23.1897
13.5759 | -34.8991 | -23.1978
13.582 | -34.8997 | -23.2064
13.5881 | -34.9002 | -23.2148
13.5942 | -34.9007 | -23.2235
13.6003 | -34.9013 | -23.232
13.6064 | -34.9018 | -23.2406
13.6125 | -34.9023 | -23.2492
13.6186 | -34.9028 | -23.2576
13.6247 | -34.9034 | -23.2662
13.6308 | -34.9039 | -23.2748
13.6369 | -34.9044 | -23.2833
13.643 | -34.9049 | -23.2919
13.6491 | -34.9054 | -23.3003
13.6552 | -34.906 | -23.3088
13.6613 | -34.9065 | -23.3173
13.6674 | -34.907 | -23.3257
13.6735 | -34.9075 | -23.3342
13.6796 | -34.9081 | -23.3427
13.6858 | -34.9086 | -23.3511
13.6919 | -34.9091 | -23.3596
13.698 | -34.9097 | -23.3682
13.7041 | -34.9102 | -23.3769
13.7102 | -34.9107 | -23.3854
13.7163 | -34.9113 | -23.3937
13.7224 | -34.9118 | -23.4023
13.7285 | -34.9123 | -23.4108
13.7346 | -34.9128 | -23.4194
13.7407 | -34.9134 | -23.4278
13.7468 | -34.9139 | -23.4363
13.7529 | -34.9144 | -23.4447
13.759 | -34.915 | -23.4532
13.7651 | -34.9155 | -23.4617
13.7712 | -34.916 | -23.4703
13.7773 | -34.9165 | -23.4789
13.7834 | -34.9171 | -23.4874
13.7895 | -34.9176 | -23.4958
13.7956 | -34.9181 | -23.5044
13.8017 | -34.9186 | -23.5128
13.8078 | -34.9192 | -23.5213
13.8139 | -34.9197 | -23.5296
13.82 | -34.9202 | -23.5379
13.8262 | -34.9207 | -23.5464
13.8323 | -34.9213 | -23.555
13.8384 | -34.9218 | -23.5634
13.8445 | -34.9224 | -23.572
13.8506 | -34.9229 | -23.5805
13.8567 | -34.9234 | -23.589
13.8628 | -34.9239 | -23.5975
13.8689 | -34.9245 | -23.606
13.875 | -34.925 | -23.6145
13.8811 | -34.9255 | -23.623
13.8872 | -34.926 | -23.6314
13.8933 | -34.9266 | -23.6396
13.8994 | -34.9271 | -23.6481
13.9055 | -34.9276 | -23.6568
13.9116 | -34.9281 | -23.6652
13.9177 | -34.9286 | -23.6737
13.9238 | -34.9292 | -23.6822
13.9299 | -34.9297 | -23.6905
13.936 | -34.9302 | -23.699
13.9421 | -34.9307 | -23.7075
13.9482 | -34.9313 | -23.7158
13.9543 | -34.9318 | -23.7244
13.9604 | -34.9323 | -23.7329
13.9665 | -34.9328 | -23.7412
13.9727 | -34.9334 | -23.7498
13.9788 | -34.9339 | -23.7583
13.9849 | -34.9345 | -23.7668
13.991 | -34.935 | -23.7753
13.9971 | -34.9355 | -23.7835
14.0032 | -34.936 | -23.7921
14.0093 | -34.9366 | -23.8005
14.0154 | -34.9371 | -23.8088
14.0215 | -34.9376 | -23.8172
14.0276 | -34.9382 | -23.8257
14.0337 | -34.9387 | -23.8341
14.0398 | -34.9393 | -23.8425
14.0459 | -34.9398 | -23.851
14.052 | -34.9403 | -23.8596
14.0581 | -34.9408 | -23.8681
14.0642 | -34.9413 | -23.8766
14.0703 | -34.9419 | -23.885
14.0764 | -34.9424 | -23.8933
14.0825 | -34.9429 | -23.9018
14.0886 | -34.9435 | -23.9102
14.0947 | -34.944 | -23.9187
14.1008 | -34.9446 | -23.9272
14.1069 | -34.9451 | -23.9354
14.1131 | -34.9456 | -23.944
14.1192 | -34.9461 | -23.9523
14.1253 | -34.9467 | -23.9606
14.1314 | -34.9472 | -23.9691
14.1375 | -34.9478 | -23.9774
14.1436 | -34.9483 | -23.9858
14.1497 | -34.9489 | -23.9942
14.1558 | -34.9494 | -24.0027
14.1619 | -34.9499 | -24.011
14.168 | -34.9504 | -24.0195
14.1741 | -34.951 | -24.0278
14.1802 | -34.9515 | -24.0363
14.1863 | -34.952 | -24.0447
14.1924 | -34.9526 | -24.0531
14.1985 | -34.9532 | -24.0614
14.2046 | -34.9537 | -24.07
14.2107 | -34.9543 | -24.0782
14.2168 | -34.9548 | -24.0867
14.2229 | -34.9553 | -24.0951
14.229 | -34.9558 | -24.1036
14.2351 | -34.9563 | -24.112
14.2412 | -34.9569 | -24.1205
14.2473 | -34.9574 | -24.1287
14.2534 | -34.958 | -24.1369
14.2596 | -34.9585 | -24.1453
14.2657 | -34.959 | -24.1537
14.2718 | -34.9595 | -24.1621
14.2779 | -34.9601 | -24.1706
14.284 | -34.9606 | -24.179
14.2901 | -34.9611 | -24.1874
14.2962 | -34.9617 | -24.1959
14.3023 | -34.9622 | -24.2043
14.3084 | -34.9627 | -24.2127
14.3145 | -34.9633 | -24.2211
14.3206 | -34.9638 | -24.2297
14.3267 | -34.9643 | -24.238
14.3328 | -34.9649 | -24.2464
14.3389 | -34.9654 | -24.2548
14.345 | -34.9659 | -24.2631
14.3511 | -34.9664 | -24.2715
14.3572 | -34.967 | -24.28
14.3633 | -34.9675 | -24.2885
14.3694 | -34.968 | -24.2969
14.3755 | -34.9686 | -24.3052
14.3816 | -34.9691 | -24.3137
14.3877 | -34.9697 | -24.322
14.3938 | -34.9702 | -24.3306
14.4 | -34.9708 | -24.3389
14.4061 | -34.9713 | -24.3472
14.4122 | -34.9718 | -24.3557
14.4183 | -34.9723 | -24.364
14.4244 | -34.9729 | -24.3724
14.4305 | -34.9734 | -24.3808
14.4366 | -34.974 | -24.3891
14.4427 | -34.9745 | -24.3976
14.4488 | -34.9751 | -24.4061
14.4549 | -34.9756 | -24.4144
14.461 | -34.9762 | -24.4227
14.4671 | -34.9767 | -24.4313
14.4732 | -34.9772 | -24.4397
14.4793 | -34.9778 | -24.4481
14.4854 | -34.9783 | -24.4564
14.4915 | -34.9788 | -24.4645
14.4976 | -34.9794 | -24.4727
14.5037 | -34.9799 | -24.4811
14.5098 | -34.9805 | -24.4894
14.5159 | -34.981 | -24.4978
14.522 | -34.9815 | -24.5062
14.5281 | -34.9821 | -24.5145
14.5342 | -34.9826 | -24.5229
14.5403 | -34.9832 | -24.5312
14.5465 | -34.9837 | -24.5396
14.5526 | -34.9842 | -24.5481
14.5587 | -34.9848 | -24.5565
14.5648 | -34.9853 | -24.5647
14.5709 | -34.9859 | -24.573
14.577 | -34.9864 | -24.5812
14.5831 | -34.9869 | -24.5897
14.5892 | -34.9875 | -24.598
14.5953 | -34.988 | -24.6062
14.6014 | -34.9885 | -24.6147
14.6075 | -34.9891 | -24.6231
14.6136 | -34.9896 | -24.6315
14.6197 | -34.9901 | -24.6399
14.6258 | -34.9907 | -24.6482
14.6319 | -34.9912 | -24.6565
14.638 | -34.9917 | -24.6648
14.6441 | -34.9923 | -24.6733
14.6502 | -34.9928 | -24.6818
14.6563 | -34.9934 | -24.6902
14.6624 | -34.9939 | -24.6986
14.6685 | -34.9945 | -24.7066
14.6746 | -34.995 | -24.715
14.6807 | -34.9955 | -24.7232
14.6869 | -34.9961 | -24.7315
14.693 | -34.9966 | -24.7398
14.6991 | -34.9972 | -24.7481
14.7052 | -34.9977 | -24.7563
14.7113 | -34.9983 | -24.7647
14.7174 | -34.9988 | -24.7731
14.7235 | -34.9994 | -24.7813
14.7296 | -34.9999 | -24.7895
14.7357 | -35.0005 | -24.7978
14.7418 | -35.001 | -24.8062
14.7479 | -35.0016 | -24.8145
14.754 | -35.0021 | -24.8229
14.7601 | -35.0027 | -24.8311
14.7662 | -35.0032 | -24.8393
14.7723 | -35.0038 | -24.8474
14.7784 | -35.0043 | -24.8558
14.7845 | -35.0048 | -24.8641
14.7906 | -35.0054 | -24.8725
14.7967 | -35.0059 | -24.8808
14.8028 | -35.0065 | -24.8891
14.8089 | -35.007 | -24.8974
14.815 | -35.0076 | -24.9057
14.8211 | -35.0081 | -24.914
14.8272 | -35.0087 | -24.9223
14.8334 | -35.0092 | -24.9306
14.8395 | -35.0098 | -24.9388
14.8456 | -35.0103 | -24.947
14.8517 | -35.0109 | -24.9554
14.8578 | -35.0114 | -24.9638
14.8639 | -35.0119 | -24.9722
14.87 | -35.0125 | -24.9805
14.8761 | -35.013 | -24.9888
14.8822 | -35.0136 | -24.9971
14.8883 | -35.0141 | -25.0055
14.8944 | -35.0146 | -25.0138
14.9005 | -35.0152 | -25.0221
14.9066 | -35.0157 | -25.0304
14.9127 | -35.0163 | -25.0385
14.9188 | -35.0168 | -25.0468
14.9249 | -35.0173 | -25.0552
14.931 | -35.0179 | -25.0635
14.9371 | -35.0185 | -25.0717
14.9432 | -35.019 | -25.08
14.9493 | -35.0196 | -25.0882
14.9554 | -35.0201 | -25.0964
14.9615 | -35.0206 | -25.1046
14.9676 | -35.0212 | -25.1131
14.9738 | -35.0217 | -25.1213
14.9799 | -35.0223 | -25.1296
14.986 | -35.0228 | -25.1378
14.9921 | -35.0233 | -25.1461
14.9982 | -35.0238 | -25.1544
15.0043 | -35.0244 | -25.1626
15.0104 | -35.0249 | -25.1711
15.0165 | -35.0255 | -25.1796
15.0226 | -35.026 | -25.1876
15.0287 | -35.0266 | -25.196
15.0348 | -35.0271 | -25.2043
15.0409 | -35.0277 | -25.2126
15.047 | -35.0282 | -25.2209
15.0531 | -35.0288 | -25.2291
15.0592 | -35.0293 | -25.2373
15.0653 | -35.0299 | -25.2456
15.0714 | -35.0304 | -25.2538
15.0775 | -35.031 | -25.262
15.0836 | -35.0315 | -25.2703
15.0897 | -35.0321 | -25.2783
15.0958 | -35.0327 | -25.2864
15.1019 | -35.0332 | -25.2948
15.108 | -35.0338 | -25.3029
15.1141 | -35.0343 | -25.3112
15.1203 | -35.0349 | -25.3194
15.1264 | -35.0354 | -25.3275
15.1325 | -35.036 | -25.3357
15.1386 | -35.0365 | -25.344
15.1447 | -35.0371 | -25.3522
15.1508 | -35.0377 | -25.3604
15.1569 | -35.0382 | -25.3687
15.163 | -35.0388 | -25.3767
15.1691 | -35.0393 | -25.3848
15.1752 | -35.0399 | -25.3931
15.1813 | -35.0404 | -25.4013
15.1874 | -35.041 | -25.4095
15.1935 | -35.0416 | -25.4175
15.1996 | -35.0421 | -25.4257
15.2057 | -35.0426 | -25.4339
15.2118 | -35.0432 | -25.4421
15.2179 | -35.0438 | -25.4504
15.224 | -35.0443 | -25.4586
15.2301 | -35.0449 | -25.4666
15.2362 | -35.0454 | -25.4748
15.2423 | -35.046 | -25.4832
15.2484 | -35.0466 | -25.4914
15.2545 | -35.0471 | -25.4997
15.2607 | -35.0477 | -25.508
15.2668 | -35.0482 | -25.5159
15.2729 | -35.0488 | -25.5241
15.279 | -35.0494 | -25.5325
15.2851 | -35.0499 | -25.5406
15.2912 | -35.0505 | -25.5489
15.2973 | -35.0511 | -25.5572
15.3034 | -35.0516 | -25.5653
15.3095 | -35.0522 | -25.5735
15.3156 | -35.0527 | -25.5817
15.3217 | -35.0533 | -25.5898
15.3278 | -35.0539 | -25.5982
15.3339 | -35.0544 | -25.6062
15.34 | -35.055 | -25.6142
15.3461 | -35.0556 | -25.6224
15.3522 | -35.0561 | -25.6306
15.3583 | -35.0566 | -25.639
15.3644 | -35.0572 | -25.6471
15.3705 | -35.0578 | -25.6551
15.3766 | -35.0583 | -25.6634
15.3827 | -35.0589 | -25.6715
15.3888 | -35.0594 | -25.6797
15.3949 | -35.06 | -25.6879
15.401 | -35.0605 | -25.6962
15.4072 | -35.0611 | -25.7042
15.4133 | -35.0616 | -25.7124
15.4194 | -35.0622 | -25.7206
15.4255 | -35.0627 | -25.7288
15.4316 | -35.0633 | -25.7369
15.4377 | -35.0638 | -25.7453
15.4438 | -35.0644 | -25.7534
15.4499 | -35.065 | -25.7615
15.456 | -35.0656 | -25.7696
15.4621 | -35.0661 | -25.7777
15.4682 | -35.0666 | -25.786
15.4743 | -35.0672 | -25.794
15.4804 | -35.0677 | -25.8022
15.4865 | -35.0683 | -25.8104
15.4926 | -35.0689 | -25.8187
15.4987 | -35.0694 | -25.8267
15.5048 | -35.07 | -25.8349
15.5109 | -35.0705 | -25.843
15.517 | -35.0711 | -25.8511
15.5231 | -35.0717 | -25.8593
15.5292 | -35.0722 | -25.8674
15.5353 | -35.0728 | -25.8757
15.5414 | -35.0733 | -25.8837
15.5476 | -35.0739 | -25.8916
15.5537 | -35.0745 | -25.8998
15.5598 | -35.075 | -25.9081
15.5659 | -35.0756 | -25.9163
15.572 | -35.0761 | -25.9244
15.5781 | -35.0767 | -25.9324
15.5842 | -35.0773 | -25.9405
15.5903 | -35.0778 | -25.9484
15.5964 | -35.0784 | -25.9564
15.6025 | -35.079 | -25.9645
15.6086 | -35.0795 | -25.9726
15.6147 | -35.0801 | -25.9809
15.6208 | -35.0807 | -25.9888
15.6269 | -35.0812 | -25.9969
15.633 | -35.0818 | -26.005
15.6391 | -35.0824 | -26.0132
15.6452 | -35.0829 | -26.0213
15.6513 | -35.0835 | -26.0291
15.6574 | -35.0841 | -26.0373
15.6635 | -35.0846 | -26.0454
15.6696 | -35.0852 | -26.0535
15.6757 | -35.0857 | -26.0617
15.6818 | -35.0863 | -26.0698
15.688 | -35.0868 | -26.0777
15.6941 | -35.0874 | -26.0857
15.7002 | -35.0879 | -26.0938
15.7063 | -35.0885 | -26.1018
15.7124 | -35.089 | -26.11
15.7185 | -35.0896 | -26.1181
15.7246 | -35.0902 | -26.1261
15.7307 | -35.0908 | -26.1341
15.7368 | -35.0913 | -26.1422
15.7429 | -35.0919 | -26.1502
15.749 | -35.0924 | -26.1583
15.7551 | -35.093 | -26.1665
15.7612 | -35.0936 | -26.1745
15.7673 | -35.0941 | -26.1825
15.7734 | -35.0947 | -26.1905
15.7795 | -35.0953 | -26.1985
15.7856 | -35.0958 | -26.2065
15.7917 | -35.0964 | -26.2144
15.7978 | -35.097 | -26.2225
15.8039 | -35.0975 | -26.2305
15.81 | -35.0981 | -26.2386
15.8161 | -35.0987 | -26.2467
15.8222 | -35.0992 | -26.2547
15.8283 | -35.0998 | -26.2625
15.8345 | -35.1004 | -26.2706
15.8406 | -35.1009 | -26.2785
15.8467 | -35.1015 | -26.2865
15.8528 | -35.102 | -26.2946
15.8589 | -35.1026 | -26.3027
15.865 | -35.1032 | -26.3104
15.8711 | -35.1038 | -26.3185
15.8772 | -35.1043 | -26.3264
15.8833 | -35.1049 | -26.3345
15.8894 | -35.1054 | -26.3427
15.8955 | -35.106 | -26.3507
15.9016 | -35.1065 | -26.3585
15.9077 | -35.1071 | -26.3665
15.9138 | -35.1076 | -26.3746
15.9199 | -35.1082 | -26.3824
15.926 | -35.1087 | -26.3904
15.9321 | -35.1093 | -26.3983
15.9382 | -35.1099 | -26.4061
15.9443 | -35.1104 | -26.4141
15.9504 | -35.111 | -26.4221
15.9565 | -35.1115 | -26.4301
15.9626 | -35.1121 | -26.4383
15.9687 | -35.1126 | -26.446
15.9749 | -35.1132 | -26.4538
15.981 | -35.1138 | -26.4618
15.9871 | -35.1143 | -26.4699
15.9932 | -35.1149 | -26.4777
15.9993 | -35.1154 | -26.4858
16.0054 | -35.116 | -26.4934
16.0115 | -35.1166 | -26.5015
16.0176 | -35.1172 | -26.5095
16.0237 | -35.1177 | -26.5176
16.0298 | -35.1183 | -26.5255
16.0359 | -35.1188 | -26.5335
16.042 | -35.1194 | -26.5414
16.0481 | -35.12 | -26.5495
16.0542 | -35.1205 | -26.5575
16.0603 | -35.1211 | -26.5654
16.0664 | -35.1216 | -26.5734
16.0725 | -35.1222 | -26.5811
16.0786 | -35.1227 | -26.5889
16.0847 | -35.1233 | -26.597
16.0908 | -35.1239 | -26.6049
16.0969 | -35.1244 | -26.6131
16.103 | -35.125 | -26.621
16.1091 | -35.1255 | -26.6289
16.1152 | -35.1261 | -26.6369
16.1214 | -35.1267 | -26.6449
16.1275 | -35.1272 | -26.6527
16.1336 | -35.1278 | -26.6606
16.1397 | -35.1284 | -26.6685
16.1458 | -35.1289 | -26.6763
16.1519 | -35.1295 | -26.6845
16.158 | -35.13 | -26.6925
16.1641 | -35.1306 | -26.7005
16.1702 | -35.1311 | -26.7085
16.1763 | -35.1317 | -26.7165
16.1824 | -35.1323 | -26.7243
16.1885 | -35.1328 | -26.7321
16.1946 | -35.1334 | -26.7399
16.2007 | -35.134 | -26.7478
16.2068 | -35.1345 | -26.7557
16.2129 | -35.1351 | -26.7637
16.219 | -35.1357 | -26.7713
16.2251 | -35.1362 | -26.7791
16.2312 | -35.1368 | -26.787
16.2373 | -35.1373 | -26.7947
16.2434 | -35.1379 | -26.8026
16.2495 | -35.1385 | -26.8104
16.2556 | -35.1391 | -26.8185
16.2618 | -35.1396 | -26.8263
16.2679 | -35.1402 | -26.8344
16.274 | -35.1407 | -26.8423
16.2801 | -35.1413 | -26.8501
16.2862 | -35.1419 | -26.8579
16.2923 | -35.1424 | -26.8658
16.2984 | -35.143 | -26.8737
16.3045 | -35.1435 | -26.8816
16.3106 | -35.1441 | -26.8895
16.3167 | -35.1447 | -26.8974
16.3228 | -35.1452 | -26.9051
16.3289 | -35.1457 | -26.9129
16.335 | -35.1463 | -26.9208
16.3411 | -35.1469 | -26.9287
16.3472 | -35.1474 | -26.9368
16.3533 | -35.148 | -26.9446
16.3594 | -35.1485 | -26.9523
16.3655 | -35.1491 | -26.9601
16.3716 | -35.1496 | -26.9679
16.3777 | -35.1502 | -26.9757
16.3838 | -35.1507 | -26.9835
16.3899 | -35.1513 | -26.9912
16.396 | -35.1519 | -26.9991
16.4021 | -35.1524 | -27.0069
16.4083 | -35.153 | -27.0146
16.4144 | -35.1535 | -27.0225
16.4205 | -35.1541 | -27.0304
16.4266 | -35.1546 | -27.038
16.4327 | -35.1552 | -27.0458
16.4388 | -35.1558 | -27.0537
16.4449 | -35.1563 | -27.0615
16.451 | -35.1569 | -27.0695
16.4571 | -35.1574 | -27.0775
16.4632 | -35.158 | -27.085
16.4693 | -35.1585 | -27.093
16.4754 | -35.1591 | -27.1009
16.4815 | -35.1596 | -27.1088
16.4876 | -35.1602 | -27.1165
16.4937 | -35.1607 | -27.1243
16.4998 | -35.1613 | -27.1318
16.5059 | -35.1619 | -27.1396
16.512 | -35.1624 | -27.1476
16.5181 | -35.163 | -27.1555
16.5242 | -35.1636 | -27.1634
16.5303 | -35.1641 | -27.1711
16.5364 | -35.1647 | -27.1788
16.5425 | -35.1652 | -27.1867
16.5487 | -35.1658 | -27.1944
16.5548 | -35.1663 | -27.2022
16.5609 | -35.1669 | -27.21
16.567 | -35.1674 | -27.2177
16.5731 | -35.168 | -27.2256
16.5792 | -35.1685 | -27.2335
16.5853 | -35.1691 | -27.2413
16.5914 | -35.1696 | -27.2491
16.5975 | -35.1702 | -27.257
16.6036 | -35.1707 | -27.2646
16.6097 | -35.1713 | -27.2725
16.6158 | -35.1718 | -27.2802
16.6219 | -35.1724 | -27.2881
16.628 | -35.1729 | -27.2959
16.6341 | -35.1735 | -27.3037
16.6402 | -35.174 | -27.3113
16.6463 | -35.1746 | -27.3192
16.6524 | -35.1751 | -27.3269
16.6585 | -35.1757 | -27.3347
16.6646 | -35.1763 | -27.3425
16.6707 | -35.1768 | -27.3503
16.6768 | -35.1774 | -27.3581
16.6829 | -35.1779 | -27.3657
16.689 | -35.1785 | -27.3734
16.6952 | -35.179 | -27.381
16.7013 | -35.1796 | -27.3887
16.7074 | -35.1801 | -27.3964
16.7135 | -35.1807 | -27.4041
16.7196 | -35.1812 | -27.4119
16.7257 | -35.1818 | -27.4196
16.7318 | -35.1824 | -27.4272
16.7379 | -35.1829 | -27.4349
16.744 | -35.1835 | -27.4425
16.7501 | -35.184 | -27.4502
16.7562 | -35.1845 | -27.4579
16.7623 | -35.1851 | -27.4654
16.7684 | -35.1856 | -27.4732
16.7745 | -35.1862 | -27.481
16.7806 | -35.1868 | -27.4884
16.7867 | -35.1874 | -27.4961
16.7928 | -35.1879 | -27.5039
16.7989 | -35.1885 | -27.5116
16.805 | -35.189 | -27.5195
16.8111 | -35.1896 | -27.5271
16.8172 | -35.1901 | -27.535
16.8233 | -35.1907 | -27.5428
16.8294 | -35.1912 | -27.5506
16.8356 | -35.1918 | -27.5584
16.8417 | -35.1923 | -27.566
16.8478 | -35.1929 | -27.5736
16.8539 | -35.1934 | -27.5812
16.86 | -35.194 | -27.5891
16.8661 | -35.1945 | -27.5967
16.8722 | -35.1951 | -27.6046
16.8783 | -35.1957 | -27.6124
16.8844 | -35.1962 | -27.6199
16.8905 | -35.1968 | -27.6276
16.8966 | -35.1973 | -27.635
16.9027 | -35.1979 | -27.6428
16.9088 | -35.1985 | -27.6507
16.9149 | -35.199 | -27.6584
16.921 | -35.1996 | -27.6659
16.9271 | -35.2001 | -27.6738
16.9332 | -35.2007 | -27.6815
16.9393 | -35.2012 | -27.6893
16.9454 | -35.2018 | -27.6968
16.9515 | -35.2023 | -27.7046
16.9576 | -35.2029 | -27.7121
16.9637 | -35.2034 | -27.7197
16.9698 | -35.204 | -27.7273
16.9759 | -35.2045 | -27.735
16.9821 | -35.2051 | -27.7427
16.9882 | -35.2057 | -27.7501
16.9943 | -35.2062 | -27.7577
17.0004 | -35.2067 | -27.7655
17.0065 | -35.2073 | -27.7732
17.0126 | -35.2079 | -27.7809
17.0187 | -35.2084 | -27.7888
17.0248 | -35.2089 | -27.7963
17.0309 | -35.2095 | -27.8039
17.037 | -35.21 | -27.8116
17.0431 | -35.2106 | -27.8192
17.0492 | -35.2111 | -27.8269
17.0553 | -35.2117 | -27.8345
17.0614 | -35.2123 | -27.8421
17.0675 | -35.2128 | -27.8497
17.0736 | -35.2134 | -27.8575
17.0797 | -35.2139 | -27.8651
17.0858 | -35.2144 | -27.8728
17.0919 | -35.215 | -27.8805
17.098 | -35.2155 | -27.8878
17.1041 | -35.2161 | -27.8955
17.1102 | -35.2167 | -27.9031
17.1163 | -35.2173 | -27.9107
17.1225 | -35.2178 | -27.9183
17.1286 | -35.2184 | -27.9258
17.1347 | -35.2189 | -27.9334
17.1408 | -35.2195 | -27.9412
17.1469 | -35.22 | -27.9487
17.153 | -35.2206 | -27.9563
17.1591 | -35.2211 | -27.9639
17.1652 | -35.2217 | -27.9713
17.1713 | -35.2222 | -27.9789
17.1774 | -35.2228 | -27.9866
17.1835 | -35.2233 | -27.9941
17.1896 | -35.2239 | -28.0017
17.1957 | -35.2244 | -28.0093
17.2018 | -35.225 | -28.0166
17.2079 | -35.2256 | -28.0243
17.214 | -35.2261 | -28.0321
17.2201 | -35.2266 | -28.0398
17.2262 | -35.2272 | -28.0475
17.2323 | -35.2277 | -28.0551
17.2384 | -35.2283 | -28.0626
17.2445 | -35.2288 | -28.0701
17.2506 | -35.2294 | -28.0778
17.2567 | -35.2299 | -28.0855
17.2628 | -35.2305 | -28.093
17.269 | -35.231 | -28.1005
17.2751 | -35.2316 | -28.1081
17.2812 | -35.2321 | -28.1159
17.2873 | -35.2326 | -28.1235
17.2934 | -35.2332 | -28.1311
17.2995 | -35.2337 | -28.1389
17.3056 | -35.2343 | -28.1462
17.3117 | -35.2349 | -28.1539
17.3178 | -35.2354 | -28.1614
17.3239 | -35.2359 | -28.1689
17.33 | -35.2365 | -28.1765
17.3361 | -35.237 | -28.1841
17.3422 | -35.2375 | -28.1915
17.3483 | -35.2381 | -28.199
17.3544 | -35.2386 | -28.2067
17.3605 | -35.2392 | -28.2142
17.3666 | -35.2397 | -28.2218
17.3727 | -35.2403 | -28.2294
17.3788 | -35.2408 | -28.2367
17.3849 | -35.2414 | -28.2443
17.391 | -35.2419 | -28.2518
17.3971 | -35.2424 | -28.2593
17.4032 | -35.243 | -28.2667
17.4094 | -35.2435 | -28.274
17.4155 | -35.2441 | -28.2815
17.4216 | -35.2446 | -28.289
17.4277 | -35.2452 | -28.2966
17.4338 | -35.2457 | -28.3041
17.4399 | -35.2463 | -28.3118
17.446 | -35.2468 | -28.319
17.4521 | -35.2474 | -28.3266
17.4582 | -35.2479 | -28.334
17.4643 | -35.2485 | -28.3416
17.4704 | -35.249 | -28.3493
17.4765 | -35.2496 | -28.3569
17.4826 | -35.2501 | -28.3643
17.4887 | -35.2507 | -28.3719
17.4948 | -35.2512 | -28.3795
17.5009 | -35.2517 | -28.387
17.507 | -35.2523 | -28.3945
17.5131 | -35.2528 | -28.4021
17.5192 | -35.2534 | -28.4094
17.5253 | -35.2539 | -28.4168
17.5314 | -35.2544 | -28.4244
17.5375 | -35.255 | -28.4319
17.5436 | -35.2555 | -28.4395
17.5497 | -35.2561 | -28.4471
17.5559 | -35.2566 | -28.4545
17.562 | -35.2572 | -28.462
17.5681 | -35.2577 | -28.4695
17.5742 | -35.2583 | -28.4771
17.5803 | -35.2588 | -28.4845
17.5864 | -35.2594 | -28.4918
17.5925 | -35.2599 | -28.4994
17.5986 | -35.2605 | -28.5068
17.6047 | -35.261 | -28.5144
17.6108 | -35.2616 | -28.5219
17.6169 | -35.2621 | -28.5292
17.623 | -35.2626 | -28.5365
17.6291 | -35.2632 | -28.5438
17.6352 | -35.2637 | -28.5513
17.6413 | -35.2643 | -28.5588
17.6474 | -35.2648 | -28.5664
17.6535 | -35.2653 | -28.5737
17.6596 | -35.2659 | -28.5811
17.6657 | -35.2665 | -28.5887
17.6718 | -35.267 | -28.5963
17.6779 | -35.2675 | -28.6038
17.684 | -35.2681 | -28.6115
17.6901 | -35.2686 | -28.619
17.6963 | -35.2691 | -28.6263
17.7024 | -35.2697 | -28.6337
17.7085 | -35.2702 | -28.641
17.7146 | -35.2708 | -28.6483
17.7207 | -35.2713 | -28.6557
17.7268 | -35.2719 | -28.663
17.7329 | -35.2724 | -28.6705
17.739 | -35.273 | -28.678
17.7451 | -35.2735 | -28.6855
17.7512 | -35.2741 | -28.6929
17.7573 | -35.2746 | -28.7005
17.7634 | -35.2751 | -28.7076
17.7695 | -35.2757 | -28.7151
17.7756 | -35.2762 | -28.7226
17.7817 | -35.2768 | -28.73
17.7878 | -35.2773 | -28.7375
17.7939 | -35.2778 | -28.7451
17.8 | -35.2784 | -28.7524
17.8061 | -35.2789 | -28.7598
17.8122 | -35.2795 | -28.7672
17.8183 | -35.28 | -28.7746
17.8244 | -35.2805 | -28.7821
17.8305 | -35.2811 | -28.7895
17.8366 | -35.2816 | -28.7967
17.8428 | -35.2822 | -28.804
17.8489 | -35.2827 | -28.8114
17.855 | -35.2832 | -28.8186
17.8611 | -35.2838 | -28.826
17.8672 | -35.2843 | -28.8332
17.8733 | -35.2849 | -28.8407
17.8794 | -35.2854 | -28.8481
17.8855 | -35.2859 | -28.8555
17.8916 | -35.2865 | -28.863
17.8977 | -35.287 | -28.8705
17.9038 | -35.2875 | -28.8774
17.9099 | -35.2881 | -28.8849
17.916 | -35.2886 | -28.8924
17.9221 | -35.2891 | -28.8998
17.9282 | -35.2897 | -28.9072
17.9343 | -35.2902 | -28.9147
17.9404 | -35.2908 | -28.922
17.9465 | -35.2913 | -28.9293
17.9526 | -35.2918 | -28.9368
17.9587 | -35.2923 | -28.9443
17.9648 | -35.2929 | -28.9517
17.9709 | -35.2934 | -28.959
17.977 | -35.294 | -28.9664
17.9832 | -35.2945 | -28.9738
17.9893 | -35.2951 | -28.9813
17.9954 | -35.2956 | -28.9886
18.0015 | -35.2961 | -28.9963
18.0076 | -35.2966 | -29.0036
18.0137 | -35.2971 | -29.0109
18.0198 | -35.2977 | -29.0182
18.0259 | -35.2982 | -29.0257
18.032 | -35.2987 | -29.0331
18.0381 | -35.2993 | -29.0404
18.0442 | -35.2998 | -29.0474
18.0503 | -35.3003 | -29.0549
18.0564 | -35.3009 | -29.0623
18.0625 | -35.3014 | -29.0699
18.0686 | -35.3019 | -29.0773
18.0747 | -35.3025 | -29.0848
18.0808 | -35.303 | -29.0919
18.0869 | -35.3036 | -29.0994
18.093 | -35.3041 | -29.1067
18.0991 | -35.3046 | -29.1141
18.1052 | -35.3052 | -29.1215
18.1113 | -35.3057 | -29.1289
18.1174 | -35.3062 | -29.136
18.1236 | -35.3068 | -29.1435
18.1297 | -35.3073 | -29.1509
18.1358 | -35.3079 | -29.1582
18.1419 | -35.3084 | -29.1655
18.148 | -35.309 | -29.1728
18.1541 | -35.3095 | -29.1799
18.1602 | -35.31 | -29.1874
18.1663 | -35.3106 | -29.1946
18.1724 | -35.3111 | -29.202
18.1785 | -35.3116 | -29.2092
18.1846 | -35.3122 | -29.2166
18.1907 | -35.3127 | -29.2239
18.1968 | -35.3132 | -29.2313
18.2029 | -35.3137 | -29.2389
18.209 | -35.3143 | -29.2462
18.2151 | -35.3148 | -29.2537
18.2212 | -35.3154 | -29.2608
18.2273 | -35.3159 | -29.2681
18.2334 | -35.3164 | -29.2755
18.2395 | -35.3169 | -29.2827
18.2456 | -35.3175 | -29.2902
18.2517 | -35.318 | -29.2975
18.2578 | -35.3185 | -29.3046
18.2639 | -35.3191 | -29.312
18.2701 | -35.3196 | -29.3193
18.2762 | -35.3202 | -29.3266
18.2823 | -35.3207 | -29.3338
18.2884 | -35.3212 | -29.3409
18.2945 | -35.3218 | -29.348
18.3006 | -35.3223 | -29.3554
18.3067 | -35.3228 | -29.3626
18.3128 | -35.3234 | -29.3701
18.3189 | -35.3239 | -29.3774
18.325 | -35.3245 | -29.3845
18.3311 | -35.325 | -29.3919
18.3372 | -35.3255 | -29.3992
18.3433 | -35.326 | -29.4066
18.3494 | -35.3266 | -29.414
18.3555 | -35.3271 | -29.4213
18.3616 | -35.3277 | -29.4284
18.3677 | -35.3282 | -29.4358
18.3738 | -35.3287 | -29.443
18.3799 | -35.3292 | -29.4504
18.386 | -35.3298 | -29.4577
18.3921 | -35.3303 | -29.465
18.3982 | -35.3308 | -29.4721
18.4043 | -35.3313 | -29.4794
18.4105 | -35.3319 | -29.4866
18.4166 | -35.3324 | -29.4939
18.4227 | -35.3329 | -29.5013
18.4288 | -35.3334 | -29.5085
18.4349 | -35.334 | -29.5157
18.441 | -35.3345 | -29.5228
18.4471 | -35.335 | -29.53
18.4532 | -35.3355 | -29.5373
18.4593 | -35.3361 | -29.5445
18.4654 | -35.3366 | -29.5514
18.4715 | -35.3371 | -29.5589
18.4776 | -35.3376 | -29.5663
18.4837 | -35.3382 | -29.5737
18.4898 | -35.3387 | -29.5811
18.4959 | -35.3392 | -29.5884
18.502 | -35.3397 | -29.5955
18.5081 | -35.3403 | -29.6028
18.5142 | -35.3408 | -29.6102
18.5203 | -35.3413 | -29.6173
18.5264 | -35.3418 | -29.6245
18.5325 | -35.3424 | -29.6318
18.5386 | -35.3429 | -29.639
18.5447 | -35.3434 | -29.6462
18.5508 | -35.344 | -29.6535
18.557 | -35.3445 | -29.6608
18.5631 | -35.345 | -29.6681
18.5692 | -35.3455 | -29.6753
18.5753 | -35.3461 | -29.6822
18.5814 | -35.3466 | -29.6894
18.5875 | -35.3471 | -29.6966
18.5936 | -35.3476 | -29.7038
18.5997 | -35.3481 | -29.711
18.6058 | -35.3486 | -29.7182
18.6119 | -35.3492 | -29.7254
18.618 | -35.3497 | -29.7328
18.6241 | -35.3502 | -29.7402
18.6302 | -35.3507 | -29.7476
18.6363 | -35.3513 | -29.7549
18.6424 | -35.3518 | -29.7621
18.6485 | -35.3523 | -29.7693
18.6546 | -35.3529 | -29.7767
18.6607 | -35.3534 | -29.784
18.6668 | -35.3539 | -29.7913
18.6729 | -35.3544 | -29.7988
18.679 | -35.355 | -29.8058
18.6851 | -35.3555 | -29.8128
18.6912 | -35.356 | -29.8198
18.6974 | -35.3566 | -29.827
18.7035 | -35.3571 | -29.8342
18.7096 | -35.3576 | -29.8416
18.7157 | -35.3582 | -29.8487
18.7218 | -35.3587 | -29.8559
18.7279 | -35.3592 | -29.8632
18.734 | -35.3598 | -29.8704
18.7401 | -35.3603 | -29.8776
18.7462 | -35.3609 | -29.8846
18.7523 | -35.3614 | -29.892
18.7584 | -35.3619 | -29.8992
18.7645 | -35.3625 | -29.9065
18.7706 | -35.363 | -29.9138
18.7767 | -35.3635 | -29.921
18.7828 | -35.3641 | -29.9278
18.7889 | -35.3646 | -29.9351
18.795 | -35.3651 | -29.9425
18.8011 | -35.3657 | -29.9498
18.8072 | -35.3662 | -29.957
18.8133 | -35.3667 | -29.9643
18.8194 | -35.3673 | -29.9714
18.8255 | -35.3678 | -29.9786
18.8316 | -35.3683 | -29.9858
18.8377 | -35.3688 | -29.993
18.8439 | -35.3694 | -30.0002
18.85 | -35.3699 | -30.0075
18.8561 | -35.3704 | -30.0146
18.8622 | -35.3709 | -30.0219
18.8683 | -35.3714 | -30.0289
18.8744 | -35.372 | -30.0362
18.8805 | -35.3725 | -30.0434
18.8866 | -35.373 | -30.0506
18.8927 | -35.3736 | -30.0577
18.8988 | -35.3741 | -30.0648
18.9049 | -35.3747 | -30.072
18.911 | -35.3752 | -30.0792
18.9171 | -35.3757 | -30.0865
18.9232 | -35.3762 | -30.0934
18.9293 | -35.3768 | -30.1006
18.9354 | -35.3773 | -30.1078
18.9415 | -35.3779 | -30.115
18.9476 | -35.3784 | -30.122
18.9537 | -35.3789 | -30.1293
18.9598 | -35.3794 | -30.1364
18.9659 | -35.3799 | -30.1435
18.972 | -35.3805 | -30.1507
18.9781 | -35.381 | -30.158
18.9843 | -35.3815 | -30.1652
18.9904 | -35.382 | -30.1724
18.9965 | -35.3825 | -30.1793
19.0026 | -35.3831 | -30.1866
19.0087 | -35.3836 | -30.1939
19.0148 | -35.3842 | -30.201
19.0209 | -35.3847 | -30.2083
19.027 | -35.3852 | -30.2155
19.0331 | -35.3858 | -30.2228
19.0392 | -35.3863 | -30.2299
19.0453 | -35.3868 | -30.237
19.0514 | -35.3873 | -30.244
19.0575 | -35.3879 | -30.2513
19.0636 | -35.3884 | -30.2579
19.0697 | -35.3889 | -30.265
19.0758 | -35.3894 | -30.2721
19.0819 | -35.3899 | -30.2793
19.088 | -35.3905 | -30.2863
19.0941 | -35.391 | -30.2935
19.1002 | -35.3915 | -30.3005
19.1063 | -35.392 | -30.3077
19.1124 | -35.3926 | -30.3149
19.1185 | -35.3931 | -30.322
19.1246 | -35.3936 | -30.3292
19.1308 | -35.3942 | -30.3364
19.1369 | -35.3947 | -30.3435
19.143 | -35.3952 | -30.3507
19.1491 | -35.3958 | -30.3579
19.1552 | -35.3963 | -30.3652
19.1613 | -35.3968 | -30.3723
19.1674 | -35.3973 | -30.3794
19.1735 | -35.3978 | -30.3865
19.1796 | -35.3983 | -30.3936
19.1857 | -35.3989 | -30.401
19.1918 | -35.3994 | -30.4081
19.1979 | -35.3999 | -30.4154
19.204 | -35.4004 | -30.4222
19.2101 | -35.401 | -30.4293
19.2162 | -35.4015 | -30.4364
19.2223 | -35.402 | -30.4436
19.2284 | -35.4025 | -30.4507
19.2345 | -35.403 | -30.4579
19.2406 | -35.4035 | -30.465
19.2467 | -35.4041 | -30.4721
19.2528 | -35.4046 | -30.4794
19.2589 | -35.4051 | -30.4865
19.265 | -35.4057 | -30.4937
19.2712 | -35.4062 | -30.5007
19.2773 | -35.4067 | -30.5077
19.2834 | -35.4072 | -30.5147
19.2895 | -35.4077 | -30.5218
19.2956 | -35.4082 | -30.5291
19.3017 | -35.4088 | -30.5361
19.3078 | -35.4093 | -30.5434
19.3139 | -35.4098 | -30.5503
19.32 | -35.4103 | -30.5574
19.3261 | -35.4108 | -30.5644
19.3322 | -35.4113 | -30.5714
19.3383 | -35.4119 | -30.5785
19.3444 | -35.4124 | -30.5854
19.3505 | -35.4129 | -30.5926
19.3566 | -35.4134 | -30.5998
19.3627 | -35.4139 | -30.6069
19.3688 | -35.4144 | -30.6139
19.3749 | -35.415 | -30.621
19.381 | -35.4155 | -30.6279
19.3871 | -35.416 | -30.6348
19.3932 | -35.4165 | -30.6422
19.3993 | -35.4171 | -30.6492
19.4054 | -35.4176 | -30.6565
19.4115 | -35.4181 | -30.6638
19.4177 | -35.4186 | -30.6708
19.4238 | -35.4192 | -30.6778
19.4299 | -35.4197 | -30.6849
19.436 | -35.4203 | -30.692
19.4421 | -35.4208 | -30.6991
19.4482 | -35.4213 | -30.7061
19.4543 | -35.4218 | -30.7128
19.4604 | -35.4223 | -30.7198
19.4665 | -35.4228 | -30.7269
19.4726 | -35.4233 | -30.7338
19.4787 | -35.4238 | -30.7408
19.4848 | -35.4243 | -30.7477
19.4909 | -35.4248 | -30.7546
19.497 | -35.4254 | -30.7617
19.5031 | -35.4259 | -30.7688
19.5092 | -35.4264 | -30.7759
19.5153 | -35.4269 | -30.7827
19.5214 | -35.4274 | -30.7896
19.5275 | -35.4279 | -30.7967
19.5336 | -35.4284 | -30.8038
19.5397 | -35.429 | -30.8109
19.5458 | -35.4295 | -30.8178
19.5519 | -35.43 | -30.8249
19.5581 | -35.4305 | -30.8317
19.5642 | -35.431 | -30.8388
19.5703 | -35.4315 | -30.8459
19.5764 | -35.432 | -30.8527
19.5825 | -35.4325 | -30.8598
19.5886 | -35.4331 | -30.867
19.5947 | -35.4336 | -30.8738
19.6008 | -35.4341 | -30.8806
19.6069 | -35.4346 | -30.8878
19.613 | -35.4351 | -30.895
19.6191 | -35.4357 | -30.9018
19.6252 | -35.4362 | -30.9087
19.6313 | -35.4367 | -30.9154
19.6374 | -35.4372 | -30.9226
19.6435 | -35.4377 | -30.9297
19.6496 | -35.4383 | -30.9366
19.6557 | -35.4388 | -30.9436
19.6618 | -35.4393 | -30.9501
19.6679 | -35.4398 | -30.9573
19.674 | -35.4403 | -30.9643
19.6801 | -35.4408 | -30.9715
19.6862 | -35.4414 | -30.9787
19.6923 | -35.4419 | -30.9855
19.6984 | -35.4424 | -30.9926
19.7046 | -35.4429 | -30.9998
19.7107 | -35.4434 | -31.0067
19.7168 | -35.4439 | -31.0137
19.7229 | -35.4444 | -31.0207
19.729 | -35.4449 | -31.0276
19.7351 | -35.4454 | -31.0346
19.7412 | -35.4459 | -31.0418
19.7473 | -35.4465 | -31.0489
19.7534 | -35.447 | -31.056
19.7595 | -35.4475 | -31.0631
19.7656 | -35.448 | -31.0701
19.7717 | -35.4485 | -31.077
19.7778 | -35.449 | -31.0841
19.7839 | -35.4495 | -31.091
19.79 | -35.45 | -31.0979
19.7961 | -35.4505 | -31.105
19.8022 | -35.451 | -31.1117
19.8083 | -35.4516 | -31.1188
19.8144 | -35.4521 | -31.1257
19.8205 | -35.4526 | -31.1326
19.8266 | -35.4531 | -31.1396
19.8327 | -35.4537 | -31.1466
19.8388 | -35.4542 | -31.1532
19.845 | -35.4547 | -31.1603
19.8511 | -35.4552 | -31.1673
19.8572 | -35.4557 | -31.1743
19.8633 | -35.4562 | -31.1813
19.8694 | -35.4566 | -31.1884
19.8755 | -35.4572 | -31.1952
19.8816 | -35.4577 | -31.2024
19.8877 | -35.4582 | -31.2094
19.8938 | -35.4587 | -31.2163
19.8999 | -35.4592 | -31.2235
19.906 | -35.4597 | -31.2306
19.9121 | -35.4602 | -31.2373
19.9182 | -35.4608 | -31.2444
19.9243 | -35.4613 | -31.2514
19.9304 | -35.4618 | -31.2583
19.9365 | -35.4623 | -31.2654
19.9426 | -35.4628 | -31.2724
19.9487 | -35.4634 | -31.2795
19.9548 | -35.4639 | -31.2865
19.9609 | -35.4644 | -31.2936
19.967 | -35.4649 | -31.3005
19.9731 | -35.4654 | -31.3074
19.9792 | -35.466 | -31.314
19.9853 | -35.4665 | -31.321
19.9915 | -35.467 | -31.3278
19.9976 | -35.4675 | -31.3347
20.0037 | -35.468 | -31.3416
20.0098 | -35.4685 | -31.3486
20.0159 | -35.469 | -31.3552
20.022 | -35.4695 | -31.3622
20.0281 | -35.47 | -31.3691
20.0342 | -35.4705 | -31.376
20.0403 | -35.4711 | -31.383
20.0464 | -35.4716 | -31.39
20.0525 | -35.4721 | -31.3968
20.0586 | -35.4726 | -31.4037
20.0647 | -35.4731 | -31.4106
20.0708 | -35.4737 | -31.4178
20.0769 | -35.4741 | -31.4249
20.083 | -35.4747 | -31.4315
20.0891 | -35.4752 | -31.4385
20.0952 | -35.4757 | -31.4455
20.1013 | -35.4762 | -31.4526
20.1074 | -35.4767 | -31.4594
20.1135 | -35.4772 | -31.4663
20.1196 | -35.4777 | -31.4729
20.1257 | -35.4783 | -31.48
20.1319 | -35.4788 | -31.487
20.138 | -35.4793 | -31.4939
20.1441 | -35.4798 | -31.5008
20.1502 | -35.4804 | -31.5078
20.1563 | -35.4809 | -31.5142
20.1624 | -35.4814 | -31.521
20.1685 | -35.4818 | -31.5279
20.1746 | -35.4824 | -31.5348
20.1807 | -35.4829 | -31.5418
20.1868 | -35.4834 | -31.5487
20.1929 | -35.4839 | -31.5554
20.199 | -35.4844 | -31.5622
20.2051 | -35.4849 | -31.569
20.2112 | -35.4855 | -31.5761
20.2173 | -35.486 | -31.583
20.2234 | -35.4865 | -31.5899
20.2295 | -35.487 | -31.5967
20.2356 | -35.4875 | -31.6035
20.2417 | -35.488 | -31.6104
20.2478 | -35.4885 | -31.6172
20.2539 | -35.489 | -31.6245
20.26 | -35.4896 | -31.6312
20.2661 | -35.4901 | -31.6382
20.2723 | -35.4907 | -31.6452
20.2784 | -35.4912 | -31.6521
20.2845 | -35.4917 | -31.6591
20.2906 | -35.4922 | -31.6657
20.2967 | -35.4927 | -31.6721
20.3028 | -35.4932 | -31.679
20.3089 | -35.4937 | -31.6859
20.315 | -35.4943 | -31.6926
20.3211 | -35.4948 | -31.6994
20.3272 | -35.4953 | -31.7064
20.3333 | -35.4958 | -31.7133
20.3394 | -35.4963 | -31.7201
20.3455 | -35.4968 | -31.7271
20.3516 | -35.4973 | -31.7337
20.3577 | -35.4978 | -31.7407
20.3638 | -35.4983 | -31.7475
20.3699 | -35.4989 | -31.7544
20.376 | -35.4994 | -31.7614
20.3821 | -35.4999 | -31.7685
20.3882 | -35.5004 | -31.7753
20.3943 | -35.5009 | -31.7821
20.4004 | -35.5014 | -31.7888
20.4065 | -35.5019 | -31.7956
20.4126 | -35.5025 | -31.8023
20.4188 | -35.5029 | -31.8091
20.4249 | -35.5035 | -31.816
20.431 | -35.504 | -31.8228
20.4371 | -35.5044 | -31.8295
20.4432 | -35.5049 | -31.8364
20.4493 | -35.5054 | -31.8433
20.4554 | -35.5059 | -31.8499
20.4615 | -35.5065 | -31.8567
20.4676 | -35.507 | -31.8638
20.4737 | -35.5075 | -31.8707
20.4798 | -35.508 | -31.8777
20.4859 | -35.5085 | -31.8846
20.492 | -35.509 | -31.8912
20.4981 | -35.5095 | -31.898
20.5042 | -35.51 | -31.9047
20.5103 | -35.5105 | -31.9115
20.5164 | -35.511 | -31.9183
20.5225 | -35.5115 | -31.9251
20.5286 | -35.5121 | -31.9318
20.5347 | -35.5125 | -31.9386
20.5408 | -35.5131 | -31.9451
20.5469 | -35.5136 | -31.9518
20.553 | -35.5141 | -31.9587
20.5592 | -35.5146 | -31.9655
20.5653 | -35.5151 | -31.9722
20.5714 | -35.5157 | -31.979
20.5775 | -35.5161 | -31.9856
20.5836 | -35.5167 | -31.9924
20.5897 | -35.5172 | -31.9994
20.5958 | -35.5177 | -32.0063
20.6019 | -35.5182 | -32.0133
20.608 | -35.5187 | -32.0202
20.6141 | -35.5192 | -32.0267
20.6202 | -35.5197 | -32.0335
20.6263 | -35.5202 | -32.0402
20.6324 | -35.5207 | -32.0471
20.6385 | -35.5212 | -32.0539
20.6446 | -35.5217 | -32.0608
20.6507 | -35.5222 | -32.0677
20.6568 | -35.5227 | -32.0742
20.6629 | -35.5232 | -32.081
20.669 | -35.5238 | -32.0879
20.6751 | -35.5243 | -32.0947
20.6812 | -35.5247 | -32.1011
20.6873 | -35.5252 | -32.1081
20.6934 | -35.5257 | -32.1148
20.6995 | -35.5263 | -32.1216
20.7057 | -35.5267 | -32.1285
20.7118 | -35.5273 | -32.1354
20.7179 | -35.5278 | -32.1417
20.724 | -35.5283 | -32.1486
20.7301 | -35.5288 | -32.1554
20.7362 | -35.5293 | -32.162
20.7423 | -35.5298 | -32.1686
20.7484 | -35.5303 | -32.1754
20.7545 | -35.5308 | -32.1823
20.7606 | -35.5313 | -32.189
20.7667 | -35.5318 | -32.1956
20.7728 | -35.5323 | -32.2026
20.7789 | -35.5328 | -32.2092
20.785 | -35.5333 | -32.216
20.7911 | -35.5338 | -32.2228
20.7972 | -35.5343 | -32.2297
20.8033 | -35.5348 | -32.2368
20.8094 | -35.5353 | -32.2435
20.8155 | -35.5358 | -32.2504
20.8216 | -35.5363 | -32.2573
20.8277 | -35.5368 | -32.264
20.8338 | -35.5373 | -32.2707
20.8399 | -35.5378 | -32.2775
20.8461 | -35.5383 | -32.2842
20.8522 | -35.5388 | -32.291
20.8583 | -35.5393 | -32.2974
20.8644 | -35.5398 | -32.3043
20.8705 | -35.5403 | -32.3111
20.8766 | -35.5409 | -32.3181
20.8827 | -35.5414 | -32.3246
20.8888 | -35.5419 | -32.3315
20.8949 | -35.5424 | -32.338
20.901 | -35.5429 | -32.3447
20.9071 | -35.5434 | -32.3515
20.9132 | -35.5439 | -32.3582
20.9193 | -35.5444 | -32.3651
20.9254 | -35.5449 | -32.3718
20.9315 | -35.5454 | -32.3785
20.9376 | -35.5459 | -32.3854
20.9437 | -35.5464 | -32.3923
20.9498 | -35.547 | -32.3988
20.9559 | -35.5475 | -32.4055
20.962 | -35.5479 | -32.4122
20.9681 | -35.5485 | -32.4186
20.9742 | -35.549 | -32.4255
20.9803 | -35.5495 | -32.4318
20.9864 | -35.55 | -32.4386
20.9926 | -35.5505 | -32.4455
20.9987 | -35.551 | -32.4519
21.0048 | -35.5516 | -32.4584
21.0109 | -35.5521 | -32.4652
21.017 | -35.5526 | -32.472
21.0231 | -35.5532 | -32.4786
21.0292 | -35.5536 | -32.4854
21.0353 | -35.5541 | -32.4918
21.0414 | -35.5547 | -32.4984
21.0475 | -35.5552 | -32.5051
21.0536 | -35.5557 | -32.5117
21.0597 | -35.5562 | -32.5181
21.0658 | -35.5567 | -32.5246
21.0719 | -35.5572 | -32.5309
21.078 | -35.5577 | -32.5376
21.0841 | -35.5582 | -32.545
21.0902 | -35.5587 | -32.5518
21.0963 | -35.5591 | -32.5588
21.1024 | -35.5596 | -32.5654
21.1085 | -35.5601 | -32.5721
21.1146 | -35.5607 | -32.5787
21.1207 | -35.5612 | -32.5854
21.1268 | -35.5617 | -32.5923
21.133 | -35.5623 | -32.599
21.1391 | -35.5628 | -32.6056
21.1452 | -35.5633 | -32.6122
21.1513 | -35.5638 | -32.6186
21.1574 | -35.5643 | -32.6255
21.1635 | -35.5648 | -32.6324
21.1696 | -35.5654 | -32.639
21.1757 | -35.5659 | -32.6454
21.1818 | -35.5664 | -32.652
21.1879 | -35.5669 | -32.6588
21.194 | -35.5674 | -32.6656
21.2001 | -35.5678 | -32.6724
21.2062 | -35.5684 | -32.6792
21.2123 | -35.5688 | -32.6853
21.2184 | -35.5693 | -32.6921
21.2245 | -35.5698 | -32.6986
21.2306 | -35.5703 | -32.7053
21.2367 | -35.5708 | -32.7122
21.2428 | -35.5713 | -32.7189
21.2489 | -35.5718 | -32.7252
21.255 | -35.5723 | -32.7322
21.2611 | -35.5728 | -32.7388
21.2672 | -35.5733 | -32.7452
21.2733 | -35.5737 | -32.7515
21.2795 | -35.5742 | -32.7582
21.2856 | -35.5747 | -32.7647
21.2917 | -35.5753 | -32.7715
21.2978 | -35.5758 | -32.7781
21.3039 | -35.5763 | -32.7847
21.31 | -35.5768 | -32.7915
21.3161 | -35.5773 | -32.7976
21.3222 | -35.5778 | -32.8041
21.3283 | -35.5783 | -32.8108
21.3344 | -35.5788 | -32.8172
21.3405 | -35.5793 | -32.8241
21.3466 | -35.5798 | -32.8307
21.3527 | -35.5803 | -32.8369
21.3588 | -35.5809 | -32.8434
21.3649 | -35.5814 | -32.85
21.371 | -35.5819 | -32.8568
21.3771 | -35.5823 | -32.8634
21.3832 | -35.5828 | -32.87
21.3893 | -35.5833 | -32.8765
21.3954 | -35.5838 | -32.8829
21.4015 | -35.5843 | -32.8894
21.4076 | -35.5848 | -32.8964
21.4137 | -35.5853 | -32.903
21.4199 | -35.5858 | -32.9098
21.426 | -35.5862 | -32.9163
21.4321 | -35.5867 | -32.9229
21.4382 | -35.5872 | -32.9297
21.4443 | -35.5877 | -32.9361
21.4504 | -35.5882 | -32.9429
21.4565 | -35.5887 | -32.9493
21.4626 | -35.5892 | -32.956
21.4687 | -35.5897 | -32.9626
21.4748 | -35.5902 | -32.9693
21.4809 | -35.5907 | -32.9757
21.487 | -35.5912 | -32.9825
21.4931 | -35.5917 | -32.9891
21.4992 | -35.5922 | -32.9957
21.5053 | -35.5927 | -33.0022
21.5114 | -35.5932 | -33.0086
21.5175 | -35.5937 | -33.0152
21.5236 | -35.5942 | -33.0217
21.5297 | -35.5947 | -33.0281
21.5358 | -35.5952 | -33.0344
21.5419 | -35.5957 | -33.041
21.548 | -35.5962 | -33.0475
21.5541 | -35.5967 | -33.0541
21.5602 | -35.5972 | -33.0607
21.5664 | -35.5977 | -33.067
21.5725 | -35.5982 | -33.0737
21.5786 | -35.5986 | -33.0805
21.5847 | -35.5991 | -33.0874
21.5908 | -35.5996 | -33.0939
21.5969 | -35.6002 | -33.1003
21.603 | -35.6007 | -33.107
21.6091 | -35.6011 | -33.1138
21.6152 | -35.6016 | -33.1204
21.6213 | -35.6021 | -33.1268
21.6274 | -35.6026 | -33.1334
21.6335 | -35.603 | -33.1398
21.6396 | -35.6035 | -33.1462
21.6457 | -35.604 | -33.1529
21.6518 | -35.6045 | -33.1594
21.6579 | -35.605 | -33.166
21.664 | -35.6055 | -33.1728
21.6701 | -35.606 | -33.1792
21.6762 | -35.6065 | -33.1856
21.6823 | -35.607 | -33.1924
21.6884 | -35.6075 | -33.1993
21.6945 | -35.6079 | -33.2057
21.7006 | -35.6084 | -33.2122
21.7068 | -35.6089 | -33.2183
21.7129 | -35.6094 | -33.2248
21.719 | -35.6098 | -33.2314
21.7251 | -35.6104 | -33.2377
21.7312 | -35.6108 | -33.2442
21.7373 | -35.6112 | -33.2503
21.7434 | -35.6118 | -33.2571
21.7495 | -35.6123 | -33.2634
21.7556 | -35.6127 | -33.2701
21.7617 | -35.6132 | -33.2768
21.7678 | -35.6137 | -33.2832
21.7739 | -35.6142 | -33.2897
21.78 | -35.6147 | -33.2962
21.7861 | -35.6152 | -33.3029
21.7922 | -35.6157 | -33.3094
21.7983 | -35.6162 | -33.316
21.8044 | -35.6166 | -33.3226
21.8105 | -35.6171 | -33.329
21.8166 | -35.6176 | -33.3357
21.8227 | -35.6182 | -33.3421
21.8288 | -35.6187 | -33.3487
21.8349 | -35.6192 | -33.3554
21.841 | -35.6196 | -33.3618
21.8471 | -35.6201 | -33.3683
21.8533 | -35.6206 | -33.3748
21.8594 | -35.6211 | -33.3813
21.8655 | -35.6216 | -33.3878
21.8716 | -35.6221 | -33.3943
21.8777 | -35.6226 | -33.4006
21.8838 | -35.6231 | -33.4072
21.8899 | -35.6236 | -33.4137
21.896 | -35.6241 | -33.4202
21.9021 | -35.6246 | -33.4267
21.9082 | -35.6251 | -33.4335
21.9143 | -35.6256 | -33.4401
21.9204 | -35.6261 | -33.4467
21.9265 | -35.6266 | -33.4531
21.9326 | -35.6271 | -33.4598
21.9387 | -35.6276 | -33.4662
21.9448 | -35.6281 | -33.4729
21.9509 | -35.6285 | -33.4791
21.957 | -35.629 | -33.4855
21.9631 | -35.6295 | -33.492
21.9692 | -35.6299 | -33.4985
21.9753 | -35.6304 | -33.5049
21.9814 | -35.6309 | -33.5114
21.9875 | -35.6313 | -33.5177
21.9937 | -35.6318 | -33.5241
21.9998 | -35.6324 | -33.5306
22.0059 | -35.6328 | -33.5373
22.012 | -35.6333 | -33.5438
22.0181 | -35.6338 | -33.5501
22.0242 | -35.6343 | -33.5563
22.0303 | -35.6348 | -33.5629
22.0364 | -35.6353 | -33.5693
22.0425 | -35.6359 | -33.5757
22.0486 | -35.6364 | -33.5822
22.0547 | -35.6368 | -33.5886
22.0608 | -35.6373 | -33.5949
22.0669 | -35.6378 | -33.6015
22.073 | -35.6382 | -33.6079
22.0791 | -35.6387 | -33.6144
22.0852 | -35.6392 | -33.6211
22.0913 | -35.6396 | -33.6276
22.0974 | -35.6401 | -33.634
22.1035 | -35.6406 | -33.6407
22.1096 | -35.641 | -33.6471
22.1157 | -35.6415 | -33.6538
22.1218 | -35.642 | -33.6603
22.1279 | -35.6425 | -33.6664
22.134 | -35.643 | -33.673
22.1402 | -35.6434 | -33.6795
22.1463 | -35.6439 | -33.6862
22.1524 | -35.6444 | -33.6928
22.1585 | -35.6449 | -33.6995
22.1646 | -35.6453 | -33.7056
22.1707 | -35.6458 | -33.7123
22.1768 | -35.6464 | -33.7184
22.1829 | -35.6468 | -33.7249
22.189 | -35.6473 | -33.7313
22.1951 | -35.6478 | -33.7375
22.2012 | -35.6483 | -33.744
22.2073 | -35.6488 | -33.7505
22.2134 | -35.6493 | -33.757
22.2195 | -35.6497 | -33.7634
22.2256 | -35.6502 | -33.7696
22.2317 | -35.6507 | -33.776
22.2378 | -35.6512 | -33.7823
22.2439 | -35.6517 | -33.7885
22.25 | -35.6522 | -33.795
22.2561 | -35.6526 | -33.8015
22.2622 | -35.6531 | -33.8078
22.2683 | -35.6536 | -33.8141
22.2744 | -35.654 | -33.8206
22.2806 | -35.6546 | -33.8271
22.2867 | -35.655 | -33.8334
22.2928 | -35.6556 | -33.8399
22.2989 | -35.6561 | -33.8462
22.305 | -35.6566 | -33.8526
22.3111 | -35.6571 | -33.859
22.3172 | -35.6576 | -33.8655
22.3233 | -35.6581 | -33.8721
22.3294 | -35.6586 | -33.8785
22.3355 | -35.659 | -33.8848
22.3416 | -35.6596 | -33.8912
22.3477 | -35.6601 | -33.8975
22.3538 | -35.6605 | -33.9039
22.3599 | -35.661 | -33.9104
22.366 | -35.6616 | -33.9168
22.3721 | -35.6621 | -33.9229
22.3782 | -35.6625 | -33.9294
22.3843 | -35.663 | -33.9359
22.3904 | -35.6635 | -33.9421
22.3965 | -35.664 | -33.9485
22.4026 | -35.6645 | -33.955
22.4087 | -35.665 | -33.961
22.4148 | -35.6655 | -33.9674
22.4209 | -35.666 | -33.9738
22.4271 | -35.6665 | -33.9802
22.4332 | -35.667 | -33.9866
22.4393 | -35.6674 | -33.993
22.4454 | -35.6679 | -33.9991
22.4515 | -35.6684 | -34.0055
22.4576 | -35.6688 | -34.0119
22.4637 | -35.6693 | -34.0182
22.4698 | -35.6698 | -34.0247
22.4759 | -35.6703 | -34.0308
22.482 | -35.6708 | -34.0371
22.4881 | -35.6713 | -34.0435
22.4942 | -35.6718 | -34.0501
22.5003 | -35.6722 | -34.0566
22.5064 | -35.6727 | -34.0631
22.5125 | -35.6732 | -34.069
22.5186 | -35.6737 | -34.0753
22.5247 | -35.6742 | -34.0819
22.5308 | -35.6747 | -34.0883
22.5369 | -35.6751 | -34.0949
22.543 | -35.6756 | -34.1014
22.5491 | -35.676 | -34.1075
22.5552 | -35.6766 | -34.1139
22.5613 | -35.6771 | -34.1204
22.5675 | -35.6776 | -34.1267
22.5736 | -35.678 | -34.1329
22.5797 | -35.6785 | -34.1393
22.5858 | -35.679 | -34.1454
22.5919 | -35.6796 | -34.1518
22.598 | -35.6801 | -34.1584
22.6041 | -35.6805 | -34.1649
22.6102 | -35.681 | -34.1712
22.6163 | -35.6815 | -34.1774
22.6224 | -35.682 | -34.1838
22.6285 | -35.6825 | -34.1899
22.6346 | -35.683 | -34.1962
22.6407 | -35.6834 | -34.2026
22.6468 | -35.6839 | -34.2091
22.6529 | -35.6844 | -34.2151
22.659 | -35.6849 | -34.2215
22.6651 | -35.6854 | -34.2278
22.6712 | -35.6858 | -34.2342
22.6773 | -35.6863 | -34.2407
22.6834 | -35.6868 | -34.2469
22.6895 | -35.6873 | -34.253
22.6956 | -35.6878 | -34.2593
22.7017 | -35.6883 | -34.2656
22.7079 | -35.6888 | -34.2718
22.714 | -35.6893 | -34.2781
22.7201 | -35.6897 | -34.2845
22.7262 | -35.6902 | -34.2907
22.7323 | -35.6907 | -34.2969
22.7384 | -35.6912 | -34.3032
22.7445 | -35.6917 | -34.3096
22.7506 | -35.6922 | -34.3159
22.7567 | -35.6926 | -34.322
22.7628 | -35.6931 | -34.3282
22.7689 | -35.6936 | -34.3345
22.775 | -35.6941 | -34.3408
22.7811 | -35.6946 | -34.3472
22.7872 | -35.6951 | -34.3534
22.7933 | -35.6956 | -34.3597
22.7994 | -35.6961 | -34.366
22.8055 | -35.6965 | -34.3724
22.8116 | -35.697 | -34.3787
22.8177 | -35.6975 | -34.3851
22.8238 | -35.698 | -34.3913
22.8299 | -35.6985 | -34.3974
22.836 | -35.699 | -34.4036
22.8421 | -35.6995 | -34.41
22.8482 | -35.6999 | -34.4164
22.8544 | -35.7004 | -34.4227
22.8605 | -35.7009 | -34.429
22.8666 | -35.7014 | -34.4349
22.8727 | -35.7019 | -34.4411
22.8788 | -35.7023 | -34.4472
22.8849 | -35.7028 | -34.4534
22.891 | -35.7033 | -34.4596
22.8971 | -35.7038 | -34.4659
22.9032 | -35.7043 | -34.4718
22.9093 | -35.7047 | -34.4781
22.9154 | -35.7052 | -34.4843
22.9215 | -35.7057 | -34.4906
22.9276 | -35.7062 | -34.4969
22.9337 | -35.7066 | -34.503
22.9398 | -35.7071 | -34.5093
22.9459 | -35.7076 | -34.5157
22.952 | -35.7081 | -34.5219
22.9581 | -35.7086 | -34.5283
22.9642 | -35.7091 | -34.5346
22.9703 | -35.7096 | -34.5407
22.9764 | -35.71 | -34.5468
22.9825 | -35.7105 | -34.5531
22.9886 | -35.711 | -34.5594
22.9948 | -35.7115 | -34.5656
23.0009 | -35.712 | -34.572
23.007 | -35.7124 | -34.578
23.0131 | -35.7129 | -34.5842
23.0192 | -35.7134 | -34.5903
23.0253 | -35.7139 | -34.5967
23.0314 | -35.7143 | -34.6031
23.0375 | -35.7148 | -34.6093
23.0436 | -35.7153 | -34.6152
23.0497 | -35.7158 | -34.6216
23.0558 | -35.7163 | -34.628
23.0619 | -35.7167 | -34.6344
23.068 | -35.7172 | -34.6406
23.0741 | -35.7177 | -34.6465
23.0802 | -35.7182 | -34.6526
23.0863 | -35.7186 | -34.6589
23.0924 | -35.7191 | -34.6652
23.0985 | -35.7196 | -34.6714
23.1046 | -35.72 | -34.6777
23.1107 | -35.7205 | -34.6838
23.1168 | -35.721 | -34.6901
23.1229 | -35.7215 | -34.6962
23.129 | -35.722 | -34.7024
23.1351 | -35.7224 | -34.7085
23.1413 | -35.7229 | -34.7148
23.1474 | -35.7234 | -34.7207
23.1535 | -35.7239 | -34.7269
23.1596 | -35.7244 | -34.7332
23.1657 | -35.7249 | -34.7395
23.1718 | -35.7253 | -34.7457
23.1779 | -35.7258 | -34.7519
23.184 | -35.7263 | -34.7578
23.1901 | -35.7268 | -34.7641
23.1962 | -35.7273 | -34.7704
23.2023 | -35.7277 | -34.7767
23.2084 | -35.7282 | -34.783
23.2145 | -35.7287 | -34.7889
23.2206 | -35.7292 | -34.7952
23.2267 | -35.7297 | -34.8013
23.2328 | -35.7301 | -34.8076
23.2389 | -35.7306 | -34.8137
23.245 | -35.7311 | -34.8198
23.2511 | -35.7316 | -34.8259
23.2572 | -35.7321 | -34.8321
23.2633 | -35.7326 | -34.8383
23.2694 | -35.733 | -34.8447
23.2755 | -35.7335 | -34.851
23.2817 | -35.734 | -34.8571
23.2878 | -35.7345 | -34.8631
23.2939 | -35.735 | -34.8693
23.3 | -35.7354 | -34.8756
23.3061 | -35.7359 | -34.8817
23.3122 | -35.7364 | -34.8879
23.3183 | -35.7369 | -34.8939
23.3244 | -35.7373 | -34.9
23.3305 | -35.7378 | -34.9061
23.3366 | -35.7383 | -34.9123
23.3427 | -35.7388 | -34.9184
23.3488 | -35.7392 | -34.9246
23.3549 | -35.7397 | -34.9306
23.361 | -35.7402 | -34.9365
23.3671 | -35.7407 | -34.9428
23.3732 | -35.7412 | -34.9488
23.3793 | -35.7416 | -34.955
23.3854 | -35.7421 | -34.9613
23.3915 | -35.7426 | -34.9673
23.3976 | -35.7431 | -34.9734
23.4037 | -35.7436 | -34.9795
23.4098 | -35.7441 | -34.9855
23.4159 | -35.7446 | -34.9916
23.422 | -35.7451 | -34.9977
23.4282 | -35.7456 | -35.0036
23.4343 | -35.7461 | -35.0098
23.4404 | -35.7465 | -35.016
23.4465 | -35.747 | -35.0221
23.4526 | -35.7475 | -35.0282
23.4587 | -35.748 | -35.0343
23.4648 | -35.7484 | -35.0401
23.4709 | -35.7489 | -35.0463
23.477 | -35.7494 | -35.0525
23.4831 | -35.7499 | -35.0586
23.4892 | -35.7504 | -35.0647
23.4953 | -35.7508 | -35.0709
23.5014 | -35.7513 | -35.0771
23.5075 | -35.7517 | -35.0833
23.5136 | -35.7522 | -35.0895
23.5197 | -35.7527 | -35.0956
23.5258 | -35.7532 | -35.1017
23.5319 | -35.7537 | -35.1075
23.538 | -35.7542 | -35.1135
23.5441 | -35.7546 | -35.1196
23.5502 | -35.7551 | -35.1258
23.5563 | -35.7556 | -35.1317
23.5624 | -35.756 | -35.1378
23.5686 | -35.7565 | -35.1438
23.5747 | -35.757 | -35.1499
23.5808 | -35.7575 | -35.1561
23.5869 | -35.7579 | -35.1623
23.593 | -35.7584 | -35.1685
23.5991 | -35.7589 | -35.1748
23.6052 | -35.7594 | -35.1804
23.6113 | -35.7598 | -35.1863
23.6174 | -35.7603 | -35.1923
23.6235 | -35.7608 | -35.1984
23.6296 | -35.7612 | -35.2046
23.6357 | -35.7617 | -35.2106
23.6418 | -35.7622 | -35.2165
23.6479 | -35.7627 | -35.2225
23.654 | -35.7632 | -35.2286
23.6601 | -35.7636 | -35.2346
23.6662 | -35.7641 | -35.2406
23.6723 | -35.7646 | -35.2464
23.6784 | -35.7651 | -35.2525
23.6845 | -35.7655 | -35.2585
23.6906 | -35.766 | -35.2645
23.6967 | -35.7665 | -35.2706
23.7028 | -35.7669 | -35.2768
23.7089 | -35.7674 | -35.2827
23.7151 | -35.7679 | -35.2888
23.7212 | -35.7684 | -35.2951
23.7273 | -35.7688 | -35.301
23.7334 | -35.7693 | -35.307
23.7395 | -35.7698 | -35.3129
23.7456 | -35.7703 | -35.3187
23.7517 | -35.7708 | -35.3246
23.7578 | -35.7712 | -35.3305
23.7639 | -35.7717 | -35.3366
23.77 | -35.7722 | -35.3428
23.7761 | -35.7727 | -35.3489
23.7822 | -35.7731 | -35.3548
23.7883 | -35.7736 | -35.3609
23.7944 | -35.7741 | -35.367
23.8005 | -35.7746 | -35.3732
23.8066 | -35.7751 | -35.3792
23.8127 | -35.7755 | -35.3853
23.8188 | -35.776 | -35.3915
23.8249 | -35.7765 | -35.3976
23.831 | -35.7769 | -35.4038
23.8371 | -35.7774 | -35.41
23.8432 | -35.7778 | -35.4159
23.8493 | -35.7783 | -35.4219
23.8555 | -35.7788 | -35.4279
23.8616 | -35.7792 | -35.434
23.8677 | -35.7797 | -35.4401
23.8738 | -35.7802 | -35.4459
23.8799 | -35.7807 | -35.452
23.886 | -35.7812 | -35.4576
23.8921 | -35.7816 | -35.4637
23.8982 | -35.7821 | -35.4695
23.9043 | -35.7825 | -35.4755
23.9104 | -35.783 | -35.4813
23.9165 | -35.7834 | -35.4872
23.9226 | -35.7839 | -35.4932
23.9287 | -35.7844 | -35.4994
23.9348 | -35.7848 | -35.5055
23.9409 | -35.7853 | -35.5114
23.947 | -35.7858 | -35.5174
23.9531 | -35.7863 | -35.5231
23.9592 | -35.7868 | -35.529
23.9653 | -35.7872 | -35.5351
23.9714 | -35.7877 | -35.541
23.9775 | -35.7881 | -35.5472
23.9836 | -35.7886 | -35.5531
23.9897 | -35.7891 | -35.559
23.9958 | -35.7896 | -35.5653
24.002 | -35.79 | -35.5711
24.0081 | -35.7905 | -35.5772
24.0142 | -35.791 | -35.5832
24.0203 | -35.7914 | -35.5893
24.0264 | -35.7919 | -35.5951
24.0325 | -35.7924 | -35.6011
24.0386 | -35.7929 | -35.6071
24.0447 | -35.7934 | -35.6131
24.0508 | -35.7938 | -35.6191
24.0569 | -35.7943 | -35.6253
24.063 | -35.7948 | -35.631
24.0691 | -35.7953 | -35.637
24.0752 | -35.7957 | -35.6432
24.0813 | -35.7962 | -35.6492
24.0874 | -35.7966 | -35.6552
24.0935 | -35.7971 | -35.6613
24.0996 | -35.7976 | -35.6671
24.1057 | -35.7981 | -35.6732
24.1118 | -35.7985 | -35.6794
24.1179 | -35.799 | -35.6855
24.124 | -35.7995 | -35.6916
24.1301 | -35.7999 | -35.6973
24.1362 | -35.8004 | -35.7032
24.1424 | -35.8009 | -35.7093
24.1485 | -35.8013 | -35.7154
24.1546 | -35.8018 | -35.7213
24.1607 | -35.8023 | -35.7274
24.1668 | -35.8028 | -35.7331
24.1729 | -35.8033 | -35.7391
24.179 | -35.8037 | -35.7448
24.1851 | -35.8042 | -35.7508
24.1912 | -35.8047 | -35.7569
24.1973 | -35.8052 | -35.763
24.2034 | -35.8057 | -35.7686
24.2095 | -35.8061 | -35.7746
24.2156 | -35.8066 | -35.7806
24.2217 | -35.8071 | -35.7865
24.2278 | -35.8075 | -35.7924
24.2339 | -35.808 | -35.7982
24.24 | -35.8085 | -35.8038
24.2461 | -35.809 | -35.8095
24.2522 | -35.8095 | -35.8155
24.2583 | -35.81 | -35.8214
24.2644 | -35.8104 | -35.8274
24.2705 | -35.8109 | -35.8331
24.2766 | -35.8114 | -35.839
24.2827 | -35.8119 | -35.8449
24.2889 | -35.8124 | -35.8511
24.295 | -35.8128 | -35.857
24.3011 | -35.8133 | -35.8628
24.3072 | -35.8138 | -35.8686
24.3133 | -35.8142 | -35.8744
24.3194 | -35.8147 | -35.8805
24.3255 | -35.8151 | -35.8865
24.3316 | -35.8156 | -35.8924
24.3377 | -35.8161 | -35.8983
24.3438 | -35.8165 | -35.9038
24.3499 | -35.817 | -35.9098
24.356 | -35.8175 | -35.9157
24.3621 | -35.8179 | -35.9217
24.3682 | -35.8184 | -35.9277
24.3743 | -35.8189 | -35.9336
24.3804 | -35.8194 | -35.9392
24.3865 | -35.8198 | -35.945
24.3926 | -35.8203 | -35.951
24.3987 | -35.8208 | -35.9569
24.4048 | -35.8213 | -35.963
24.4109 | -35.8217 | -35.9688
24.417 | -35.8222 | -35.9747
24.4231 | -35.8227 | -35.9805
24.4293 | -35.8232 | -35.9865
24.4354 | -35.8236 | -35.9923
24.4415 | -35.8241 | -35.9983
24.4476 | -35.8246 | -36.0041
24.4537 | -35.8251 | -36.01
24.4598 | -35.8255 | -36.0158
24.4659 | -35.826 | -36.0216
24.472 | -35.8265 | -36.0274
24.4781 | -35.827 | -36.0332
24.4842 | -35.8275 | -36.039
24.4903 | -35.8279 | -36.0451
24.4964 | -35.8284 | -36.0509
24.5025 | -35.8288 | -36.0568
24.5086 | -35.8293 | -36.0626
24.5147 | -35.8298 | -36.0686
24.5208 | -35.8303 | -36.0742
24.5269 | -35.8308 | -36.0801
24.533 | -35.8312 | -36.086
24.5391 | -35.8317 | -36.0918
24.5452 | -35.8322 | -36.0978
24.5513 | -35.8327 | -36.1033
24.5574 | -35.8332 | -36.1091
24.5635 | -35.8337 | -36.1147
24.5696 | -35.8341 | -36.1206
24.5758 | -35.8346 | -36.1262
24.5819 | -35.835 | -36.1319
24.588 | -35.8354 | -36.1377
24.5941 | -35.8359 | -36.1436
24.6002 | -35.8363 | -36.1495
24.6063 | -35.8368 | -36.1555
24.6124 | -35.8373 | -36.1613
24.6185 | -35.8378 | -36.1672
24.6246 | -35.8382 | -36.1727
24.6307 | -35.8387 | -36.1788
24.6368 | -35.8392 | -36.1848
24.6429 | -35.8396 | -36.1907
24.649 | -35.8401 | -36.1968
24.6551 | -35.8406 | -36.2027
24.6612 | -35.841 | -36.2082
24.6673 | -35.8415 | -36.2141
24.6734 | -35.8419 | -36.22
24.6795 | -35.8424 | -36.2259
24.6856 | -35.8429 | -36.2317
24.6917 | -35.8434 | -36.2377
24.6978 | -35.8438 | -36.2432
24.7039 | -35.8443 | -36.2489
24.71 | -35.8447 | -36.2549
24.7162 | -35.8452 | -36.261
24.7223 | -35.8457 | -36.2668
24.7284 | -35.8462 | -36.2723
24.7345 | -35.8466 | -36.2779
24.7406 | -35.8471 | -36.2838
24.7467 | -35.8475 | -36.2897
24.7528 | -35.848 | -36.2955
24.7589 | -35.8485 | -36.3013
24.765 | -35.8489 | -36.3068
24.7711 | -35.8494 | -36.3128
24.7772 | -35.8499 | -36.3185
24.7833 | -35.8503 | -36.3243
24.7894 | -35.8507 | -36.33
24.7955 | -35.8512 | -36.3358
24.8016 | -35.8517 | -36.3413
24.8077 | -35.8521 | -36.3471
24.8138 | -35.8526 | -36.3529
24.8199 | -35.8531 | -36.3587
24.826 | -35.8535 | -36.3645
24.8321 | -35.854 | -36.3702
24.8382 | -35.8545 | -36.3759
24.8443 | -35.8549 | -36.3817
24.8504 | -35.8554 | -36.3876
24.8565 | -35.8558 | -36.3934
24.8627 | -35.8563 | -36.3992
24.8688 | -35.8567 | -36.4051
24.8749 | -35.8572 | -36.4109
24.881 | -35.8577 | -36.4167
24.8871 | -35.8581 | -36.4225
24.8932 | -35.8586 | -36.4283
24.8993 | -35.8591 | -36.4343
24.9054 | -35.8595 | -36.4401
24.9115 | -35.86 | -36.446
24.9176 | -35.8604 | -36.4516
24.9237 | -35.8609 | -36.4577
24.9298 | -35.8614 | -36.4635
24.9359 | -35.8619 | -36.4694
24.942 | -35.8623 | -36.4751
24.9481 | -35.8628 | -36.4809
24.9542 | -35.8633 | -36.4867
24.9603 | -35.8637 | -36.4924
24.9664 | -35.8642 | -36.4982
24.9725 | -35.8647 | -36.5044
24.9786 | -35.8652 | -36.5092
24.9847 | -35.8656 | -36.5148
24.9908 | -35.8661 | -36.5205
24.9969 | -35.8665 | -36.5265
25.0031 | -35.867 | -36.5322
25.0092 | -35.8674 | -36.5379
25.0153 | -35.8679 | -36.5436
25.0214 | -35.8684 | -36.5496
25.0275 | -35.8689 | -36.5554
25.0336 | -35.8693 | -36.5611
25.0397 | -35.8698 | -36.567
25.0458 | -35.8703 | -36.5723
25.0519 | -35.8707 | -36.578
25.058 | -35.8712 | -36.5836
25.0641 | -35.8717 | -36.5893
25.0702 | -35.8721 | -36.595
25.0763 | -35.8726 | -36.6006
25.0824 | -35.8731 | -36.6062
25.0885 | -35.8735 | -36.6119
25.0946 | -35.874 | -36.6177
25.1007 | -35.8744 | -36.6234
25.1068 | -35.8749 | -36.6291
25.1129 | -35.8754 | -36.6349
25.119 | -35.8758 | -36.6405
25.1251 | -35.8763 | -36.6461
25.1312 | -35.8768 | -36.6517
25.1373 | -35.8772 | -36.6575
25.1435 | -35.8777 | -36.6634
25.1496 | -35.8782 | -36.6688
25.1557 | -35.8786 | -36.6746
25.1618 | -35.8791 | -36.6805
25.1679 | -35.8795 | -36.6864
25.174 | -35.88 | -36.6921
25.1801 | -35.8805 | -36.698
25.1862 | -35.8809 | -36.7037
25.1923 | -35.8813 | -36.7094
25.1984 | -35.8818 | -36.7147
25.2045 | -35.8823 | -36.7204
25.2106 | -35.8827 | -36.726
25.2167 | -35.8832 | -36.7318
25.2228 | -35.8836 | -36.7374
25.2289 | -35.8841 | -36.7428
25.235 | -35.8846 | -36.7486
25.2411 | -35.8851 | -36.7543
25.2472 | -35.8855 | -36.7602
25.2533 | -35.8859 | -36.766
25.2594 | -35.8864 | -36.7715
25.2655 | -35.8869 | -36.7773
25.2716 | -35.8873 | -36.7828
25.2777 | -35.8878 | -36.7886
25.2838 | -35.8882 | -36.7944
25.29 | -35.8887 | -36.8
25.2961 | -35.8892 | -36.8056
25.3022 | -35.8897 | -36.8114
25.3083 | -35.8901 | -36.8172
25.3144 | -35.8906 | -36.8229
25.3205 | -35.8911 | -36.8285
25.3266 | -35.8915 | -36.834
25.3327 | -35.892 | -36.8398
25.3388 | -35.8924 | -36.8456
25.3449 | -35.8929 | -36.851
25.351 | -35.8934 | -36.8565
25.3571 | -35.8938 | -36.8621
25.3632 | -35.8943 | -36.8679
25.3693 | -35.8948 | -36.8737
25.3754 | -35.8952 | -36.8792
25.3815 | -35.8957 | -36.8851
25.3876 | -35.8961 | -36.8908
25.3937 | -35.8966 | -36.8967
25.3998 | -35.897 | -36.9021
25.4059 | -35.8975 | -36.9078
25.412 | -35.8979 | -36.9134
25.4181 | -35.8984 | -36.9191
25.4242 | -35.8989 | -36.9249
25.4304 | -35.8993 | -36.9305
25.4365 | -35.8998 | -36.9359
25.4426 | -35.9002 | -36.9417
25.4487 | -35.9007 | -36.9474
25.4548 | -35.9012 | -36.953
25.4609 | -35.9016 | -36.9586
25.467 | -35.9021 | -36.964
25.4731 | -35.9026 | -36.9696
25.4792 | -35.903 | -36.9753
25.4853 | -35.9035 | -36.981
25.4914 | -35.9039 | -36.9866
25.4975 | -35.9044 | -36.992
25.5036 | -35.9048 | -36.9976
25.5097 | -35.9053 | -37.0033
25.5158 | -35.9058 | -37.0089
25.5219 | -35.9062 | -37.0148
25.528 | -35.9067 | -37.0204
25.5341 | -35.9072 | -37.026
25.5402 | -35.9077 | -37.0315
25.5463 | -35.9082 | -37.037
25.5524 | -35.9086 | -37.0427
25.5585 | -35.9091 | -37.048
25.5646 | -35.9096 | -37.0539
25.5707 | -35.91 | -37.0596
25.5769 | -35.9104 | -37.065
25.583 | -35.9109 | -37.0708
25.5891 | -35.9113 | -37.0762
25.5952 | -35.9118 | -37.0819
25.6013 | -35.9122 | -37.0877
25.6074 | -35.9127 | -37.0933
25.6135 | -35.9132 | -37.099
25.6196 | -35.9136 | -37.1047
25.6257 | -35.914 | -37.1104
25.6318 | -35.9145 | -37.116
25.6379 | -35.915 | -37.1217
25.644 | -35.9154 | -37.1273
25.6501 | -35.9158 | -37.1328
25.6562 | -35.9163 | -37.1383
25.6623 | -35.9168 | -37.1438
25.6684 | -35.9173 | -37.1494
25.6745 | -35.9177 | -37.1552
25.6806 | -35.9181 | -37.1605
25.6867 | -35.9186 | -37.1663
25.6928 | -35.9191 | -37.1722
25.6989 | -35.9195 | -37.1777
25.705 | -35.9199 | -37.1835
25.7111 | -35.9204 | -37.1888
25.7173 | -35.9209 | -37.1943
25.7234 | -35.9213 | -37.2
25.7295 | -35.9218 | -37.2058
25.7356 | -35.9222 | -37.2116
25.7417 | -35.9227 | -37.2171
25.7478 | -35.923 | -37.2228
25.7539 | -35.9236 | -37.2285
25.76 | -35.924 | -37.2337
25.7661 | -35.9245 | -37.2395
25.7722 | -35.9249 | -37.2452
25.7783 | -35.9254 | -37.2509
25.7844 | -35.9259 | -37.2563
25.7905 | -35.9263 | -37.262
25.7966 | -35.9267 | -37.2676
25.8027 | -35.9272 | -37.2731
25.8088 | -35.9276 | -37.2789
25.8149 | -35.928 | -37.2845
25.821 | -35.9285 | -37.2902
25.8271 | -35.9289 | -37.2958
25.8332 | -35.9294 | -37.3014
25.8393 | -35.9298 | -37.307
25.8454 | -35.9303 | -37.3127
25.8515 | -35.9307 | -37.3185
25.8576 | -35.9311 | -37.3238
25.8638 | -35.9316 | -37.3292
25.8699 | -35.932 | -37.335
25.876 | -35.9325 | -37.3405
25.8821 | -35.9329 | -37.3459
25.8882 | -35.9334 | -37.3513
25.8943 | -35.9338 | -37.3571
25.9004 | -35.9342 | -37.3627
25.9065 | -35.9347 | -37.3685
25.9126 | -35.9352 | -37.3741
25.9187 | -35.9356 | -37.3796
25.9248 | -35.9361 | -37.3848
25.9309 | -35.9365 | -37.3903
25.937 | -35.9369 | -37.3959
25.9431 | -35.9374 | -37.4015
25.9492 | -35.9378 | -37.407
25.9553 | -35.9383 | -37.4129
25.9614 | -35.9387 | -37.4183
25.9675 | -35.9391 | -37.424
25.9736 | -35.9396 | -37.4297
25.9797 | -35.94 | -37.4354
25.9858 | -35.9405 | -37.4411
25.9919 | -35.9409 | -37.447
25.998 | -35.9414 | -37.4526
26.0042 | -35.9418 | -37.4581
26.0103 | -35.9422 | -37.4638
26.0164 | -35.9426 | -37.4692
26.0225 | -35.9431 | -37.4748
26.0286 | -35.9435 | -37.4802
26.0347 | -35.944 | -37.4853
26.0408 | -35.9444 | -37.4908
26.0469 | -35.9449 | -37.4961
26.053 | -35.9453 | -37.5021
26.0591 | -35.9457 | -37.5075
26.0652 | -35.9461 | -37.5127
26.0713 | -35.9465 | -37.5182
26.0774 | -35.947 | -37.5236
26.0835 | -35.9474 | -37.5294
26.0896 | -35.9478 | -37.5349
26.0957 | -35.9483 | -37.5408
26.1018 | -35.9487 | -37.5463
26.1079 | -35.9491 | -37.5518
26.114 | -35.9496 | -37.5573
26.1201 | -35.95 | -37.5629
26.1262 | -35.9505 | -37.5687
26.1323 | -35.9509 | -37.5741
26.1384 | -35.9512 | -37.5796
26.1445 | -35.9517 | -37.5848
26.1507 | -35.9523 | -37.5907
26.1568 | -35.9527 | -37.5961
26.1629 | -35.9532 | -37.6016
26.169 | -35.9536 | -37.6071
26.1751 | -35.954 | -37.6127
26.1812 | -35.9544 | -37.6182
26.1873 | -35.9549 | -37.6239
26.1934 | -35.9553 | -37.6293
26.1995 | -35.9558 | -37.6347
26.2056 | -35.9562 | -37.6397
26.2117 | -35.9566 | -37.6452
26.2178 | -35.9571 | -37.6507
26.2239 | -35.9575 | -37.6562
26.23 | -35.9579 | -37.6616
26.2361 | -35.9584 | -37.6674
26.2422 | -35.9588 | -37.6729
26.2483 | -35.9592 | -37.6786
26.2544 | -35.9596 | -37.6841
26.2605 | -35.9601 | -37.6897
26.2666 | -35.9605 | -37.6952
26.2727 | -35.961 | -37.7008
26.2788 | -35.9614 | -37.7061
26.2849 | -35.9619 | -37.7118
26.2911 | -35.9623 | -37.7173
26.2972 | -35.9628 | -37.7229
26.3033 | -35.9632 | -37.7281
26.3094 | -35.9637 | -37.7335
26.3155 | -35.9641 | -37.739
26.3216 | -35.9645 | -37.7446
26.3277 | -35.9649 | -37.7502
26.3338 | -35.9653 | -37.7558
26.3399 | -35.9657 | -37.7608
26.346 | -35.9663 | -37.7662
26.3521 | -35.9667 | -37.7719
26.3582 | -35.9671 | -37.7773
26.3643 | -35.9676 | -37.783
26.3704 | -35.968 | -37.7885
26.3765 | -35.9685 | -37.7941
26.3826 | -35.9688 | -37.7997
26.3887 | -35.9693 | -37.8053
26.3948 | -35.9697 | -37.8108
26.4009 | -35.9701 | -37.8162
26.407 | -35.9705 | -37.8217
26.4131 | -35.971 | -37.8274
26.4192 | -35.9714 | -37.8325
26.4253 | -35.9718 | -37.8381
26.4314 | -35.9723 | -37.8433
26.4376 | -35.9727 | -37.8487
26.4437 | -35.9731 | -37.8545
26.4498 | -35.9736 | -37.8601
26.4559 | -35.974 | -37.8654
26.462 | -35.9744 | -37.8708
26.4681 | -35.9748 | -37.876
26.4742 | -35.9753 | -37.8816
26.4803 | -35.9757 | -37.8873
26.4864 | -35.9761 | -37.8931
26.4925 | -35.9766 | -37.8987
26.4986 | -35.977 | -37.9039
26.5047 | -35.9774 | -37.9093
26.5108 | -35.9779 | -37.9148
26.5169 | -35.9783 | -37.9204
26.523 | -35.9787 | -37.9258
26.5291 | -35.9792 | -37.9312
26.5352 | -35.9796 | -37.9367
26.5413 | -35.98 | -37.9422
26.5474 | -35.9805 | -37.9477
26.5535 | -35.9809 | -37.9533
26.5596 | -35.9813 | -37.9586
26.5657 | -35.9817 | -37.9641
26.5718 | -35.9822 | -37.9696
26.578 | -35.9826 | -37.975
26.5841 | -35.9831 | -37.9806
26.5902 | -35.9835 | -37.9857
26.5963 | -35.9839 | -37.9909
26.6024 | -35.9843 | -37.9962
26.6085 | -35.9848 | -38.0018
26.6146 | -35.9852 | -38.0073
26.6207 | -35.9856 | -38.0129
26.6268 | -35.986 | -38.0184
26.6329 | -35.9863 | -38.0238
26.639 | -35.9867 | -38.0291
26.6451 | -35.9871 | -38.0347
26.6512 | -35.9875 | -38.0401
26.6573 | -35.9879 | -38.0457
26.6634 | -35.9884 | -38.0511
26.6695 | -35.9887 | -38.0566
26.6756 | -35.9891 | -38.0626
26.6817 | -35.9896 | -38.0682
26.6878 | -35.99 | -38.0739
26.6939 | -35.9903 | -38.0795
26.7 | -35.9907 | -38.0852
26.7061 | -35.9911 | -38.0903
26.7122 | -35.9916 | -38.0955
26.7183 | -35.992 | -38.1011
26.7245 | -35.9924 | -38.1064
26.7306 | -35.9928 | -38.1116
26.7367 | -35.9932 | -38.1172
26.7428 | -35.9936 | -38.123
26.7489 | -35.994 | -38.1288
26.755 | -35.9945 | -38.134
26.7611 | -35.9949 | -38.1396
26.7672 | -35.9953 | -38.1449
26.7733 | -35.9957 | -38.1506
26.7794 | -35.9962 | -38.1559
26.7855 | -35.9966 | -38.1614
26.7916 | -35.9971 | -38.1667
26.7977 | -35.9975 | -38.1723
26.8038 | -35.9979 | -38.1776
26.8099 | -35.9983 | -38.1831
26.816 | -35.9988 | -38.1883
26.8221 | -35.9992 | -38.194
26.8282 | -35.9996 | -38.1993
26.8343 | -36.0001 | -38.2047
26.8404 | -36.0005 | -38.2101
26.8465 | -36.0009 | -38.2154
26.8526 | -36.0013 | -38.2208
26.8587 | -36.0017 | -38.2264
26.8649 | -36.0021 | -38.2318
26.871 | -36.0025 | -38.2372
26.8771 | -36.003 | -38.2424
26.8832 | -36.0034 | -38.2479
26.8893 | -36.0038 | -38.2534
26.8954 | -36.0042 | -38.2589
26.9015 | -36.0047 | -38.2643
26.9076 | -36.0051 | -38.2696
26.9137 | -36.0054 | -38.2754
26.9198 | -36.0058 | -38.2809
26.9259 | -36.0063 | -38.2858
26.932 | -36.0068 | -38.2916
26.9381 | -36.0072 | -38.297
26.9442 | -36.0076 | -38.3023
26.9503 | -36.0081 | -38.3078
26.9564 | -36.0085 | -38.3131
26.9625 | -36.0089 | -38.3183
26.9686 | -36.0094 | -38.3237
26.9747 | -36.0098 | -38.3288
26.9808 | -36.0102 | -38.3345
26.9869 | -36.0106 | -38.3399
26.993 | -36.011 | -38.3456
26.9991 | -36.0114 | -38.3505
27.0052 | -36.0119 | -38.3558
27.0114 | -36.0123 | -38.3612
27.0175 | -36.0127 | -38.3667
27.0236 | -36.0131 | -38.3721
27.0297 | -36.0136 | -38.3777
27.0358 | -36.014 | -38.383
27.0419 | -36.0144 | -38.3883
27.048 | -36.0148 | -38.394
27.0541 | -36.0153 | -38.3994
27.0602 | -36.0157 | -38.4048
27.0663 | -36.0161 | -38.4102
27.0724 | -36.0165 | -38.4155
27.0785 | -36.0169 | -38.4211
27.0846 | -36.0173 | -38.4264
27.0907 | -36.0178 | -38.4319
27.0968 | -36.0182 | -38.4374
27.1029 | -36.0186 | -38.4428
27.109 | -36.0191 | -38.448
27.1151 | -36.0195 | -38.4535
27.1212 | -36.0197 | -38.4586
27.1273 | -36.0202 | -38.4643
27.1334 | -36.0206 | -38.4697
27.1395 | -36.021 | -38.4752
27.1456 | -36.0215 | -38.4805
27.1518 | -36.0219 | -38.4859
27.1579 | -36.0223 | -38.4912
27.164 | -36.0227 | -38.4964
27.1701 | -36.023 | -38.5017
27.1762 | -36.0235 | -38.507
27.1823 | -36.0239 | -38.5121
27.1884 | -36.0243 | -38.5174
27.1945 | -36.0246 | -38.5228
27.2006 | -36.025 | -38.5283
27.2067 | -36.0254 | -38.5335
27.2128 | -36.0258 | -38.539
27.2189 | -36.0262 | -38.5443
27.225 | -36.0266 | -38.5499
27.2311 | -36.027 | -38.555
27.2372 | -36.0274 | -38.5602
27.2433 | -36.0279 | -38.5655
27.2494 | -36.0283 | -38.5708
27.2555 | -36.0287 | -38.5763
27.2616 | -36.0292 | -38.5819
27.2677 | -36.0296 | -38.5874
27.2738 | -36.03 | -38.5929
27.2799 | -36.0304 | -38.5983
27.286 | -36.0308 | -38.6037
27.2921 | -36.0312 | -38.6091
27.2983 | -36.0316 | -38.6146
27.3044 | -36.032 | -38.6199
27.3105 | -36.0324 | -38.6255
27.3166 | -36.0328 | -38.631
27.3227 | -36.0332 | -38.6366
27.3288 | -36.0336 | -38.642
27.3349 | -36.0339 | -38.6473
27.341 | -36.0344 | -38.6525
27.3471 | -36.0348 | -38.6579
27.3532 | -36.0351 | -38.6632
27.3593 | -36.0356 | -38.6684
27.3654 | -36.0359 | -38.6738
27.3715 | -36.0363 | -38.6789
27.3776 | -36.0367 | -38.6843
27.3837 | -36.0372 | -38.6898
27.3898 | -36.0376 | -38.6953
27.3959 | -36.038 | -38.7005
27.402 | -36.0384 | -38.7055
27.4081 | -36.0388 | -38.711
27.4142 | -36.0392 | -38.7164
27.4203 | -36.0396 | -38.7219
27.4264 | -36.04 | -38.7273
27.4325 | -36.0405 | -38.7328
27.4387 | -36.0408 | -38.7381
27.4448 | -36.0412 | -38.7434
27.4509 | -36.0416 | -38.7488
27.457 | -36.042 | -38.7541
27.4631 | -36.0424 | -38.7594
27.4692 | -36.0428 | -38.7645
27.4753 | -36.0432 | -38.7698
27.4814 | -36.0436 | -38.7754
27.4875 | -36.044 | -38.7808
27.4936 | -36.0444 | -38.7864
27.4997 | -36.0448 | -38.7918
27.5058 | -36.0451 | -38.7976
27.5119 | -36.0455 | -38.8031
27.518 | -36.046 | -38.8086
27.5241 | -36.0464 | -38.8138
27.5302 | -36.0468 | -38.8194
27.5363 | -36.0472 | -38.825
27.5424 | -36.0476 | -38.8303
27.5485 | -36.0481 | -38.8355
27.5546 | -36.0485 | -38.8406
27.5607 | -36.0488 | -38.8459
27.5668 | -36.0492 | -38.8512
27.5729 | -36.0496 | -38.8568
27.5791 | -36.05 | -38.8615
27.5852 | -36.0504 | -38.8666
27.5913 | -36.0509 | -38.872
27.5974 | -36.0512 | -38.8774
27.6035 | -36.0516 | -38.883
27.6096 | -36.052 | -38.8885
27.6157 | -36.0523 | -38.8937
27.6218 | -36.0527 | -38.8991
27.6279 | -36.0531 | -38.9047
27.634 | -36.0535 | -38.9102
27.6401 | -36.0539 | -38.9158
27.6462 | -36.0544 | -38.9212
27.6523 | -36.0548 | -38.9268
27.6584 | -36.0552 | -38.9321
27.6645 | -36.0556 | -38.9376
27.6706 | -36.056 | -38.9432
27.6767 | -36.0564 | -38.9486
27.6828 | -36.0567 | -38.9542
27.6889 | -36.0571 | -38.9598
27.695 | -36.0576 | -38.9648
27.7011 | -36.0579 | -38.9704
27.7072 | -36.0583 | -38.9755
27.7133 | -36.0587 | -38.981
27.7194 | -36.0591 | -38.9867
27.7256 | -36.0595 | -38.9922
27.7317 | -36.0599 | -38.9974
27.7378 | -36.0603 | -39.0032
27.7439 | -36.0606 | -39.0088
27.75 | -36.061 | -39.0142
27.7561 | -36.0614 | -39.0195
27.7622 | -36.0618 | -39.0249
27.7683 | -36.0622 | -39.0304
27.7744 | -36.0627 | -39.0356
27.7805 | -36.0631 | -39.0413
27.7866 | -36.0635 | -39.0471
27.7927 | -36.0639 | -39.0524
27.7988 | -36.0643 | -39.0575
27.8049 | -36.0647 | -39.063
27.811 | -36.065 | -39.0685
27.8171 | -36.0654 | -39.074
27.8232 | -36.0658 | -39.0793
27.8293 | -36.0662 | -39.0846
27.8354 | -36.0666 | -39.0902
27.8415 | -36.067 | -39.0954
27.8476 | -36.0673 | -39.101
27.8537 | -36.0677 | -39.106
27.8598 | -36.0681 | -39.111
27.866 | -36.0685 | -39.1165
27.8721 | -36.0688 | -39.1219
27.8782 | -36.0693 | -39.1274
27.8843 | -36.0696 | -39.1329
27.8904 | -36.07 | -39.1385
27.8965 | -36.0704 | -39.1435
27.9026 | -36.0708 | -39.1491
27.9087 | -36.0712 | -39.1541
27.9148 | -36.0716 | -39.1597
27.9209 | -36.0719 | -39.1653
27.927 | -36.0724 | -39.1706
27.9331 | -36.0727 | -39.1753
27.9392 | -36.0731 | -39.1811
27.9453 | -36.0735 | -39.1862
27.9514 | -36.0739 | -39.1915
27.9575 | -36.0743 | -39.1968
27.9636 | -36.0747 | -39.2019
27.9697 | -36.0751 | -39.2069
27.9758 | -36.0754 | -39.2124
27.9819 | -36.0758 | -39.2179
27.988 | -36.0762 | -39.2232
27.9941 | -36.0766 | -39.2287
28.0002 | -36.0771 | -39.2337
28.0063 | -36.0774 | -39.2393
28.0125 | -36.0778 | -39.2444
28.0186 | -36.0783 | -39.2499
28.0247 | -36.0787 | -39.2553
28.0308 | -36.0791 | -39.2607
28.0369 | -36.0795 | -39.266
28.043 | -36.0798 | -39.2718
28.0491 | -36.0802 | -39.2766
28.0552 | -36.0805 | -39.2821
28.0613 | -36.0809 | -39.2873
28.0674 | -36.0813 | -39.2927
28.0735 | -36.0817 | -39.2981

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.487 | -41.9093
21.4908 | -35.5305 | -34.8526
16.4948 | -34.9224 | -27.5028
12.6602 | -34.5621 | -21.2485
9.7171 | -34.3495 | -16.3725
7.4582 | -34.203 | -12.4133
5.7244 | -34.1003 | -9.2019
4.3936 | -34.032 | -6.6112
3.3722 | -34.0079 | -4.5684
2.5883 | -33.9953 | -3.022
1.9866 | -33.9928 | -1.7863
1.5248 | -34.0017 | -0.79683
1.1703 | -34.0042 | -0.15073
1 | -34.0137 | 0.16859
0.89822 | -34.007 | 0.43644
0.68941 | -34.0019 | 1.0132
0.52919 | -33.9998 | 1.4627
0.40612 | -33.9927 | 2.0694
0.31171 | -33.9876 | 2.5182
0.23928 | -33.995 | 3.2444
0.18358 | -34.0022 | 3.6596
0.14091 | -34.0167 | 4.5435
0.10819 | -34.0365 | 5.7035
0.083017 | -34.0695 | 7.1745
0.063705 | -34.105 | 9.0494
0.048923 | -34.1834 | 11.5406
0.037479 | -34.2878 | 14.7345
0.028801 | -34.4713 | 18.7349
0.022078 | -34.776 | 24.0621
0.017023 | -35.255 | 30.4686
0.013018 | -36.0675 | 38.1605
0.010014 | -37.1677 | 46.0605

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003704 | -10.0448
0.007407 | -27.7458
0.011111 | -35.6833
0.014815 | -41.528
0.018519 | -44.9497
0.022222 | -47.094
0.025926 | -47.9434
0.02963 | -49.2649
0.033333 | -51.1433
0.037037 | -52.6447
0.040741 | -54.4785
0.044444 | -55.1318
0.048148 | -54.8597
0.051852 | -56.269
0.055556 | -57.6532
0.059259 | -58.365
0.062963 | -58.2523
0.066667 | -59.1828
0.07037 | -59.8665
0.074074 | -60.514
0.077778 | -60.8775
0.081481 | -61.9295
0.085185 | -62.4556
0.088889 | -62.7999
0.092593 | -64.3896
0.096296 | -65.1183
0.1 | -64.7221
0.1037 | -64.6182
0.10741 | -64.8679
0.11111 | -64.682
0.11482 | -65.2244
0.11852 | -66.4094
0.12222 | -66.9486
0.12593 | -67.3549
0.12963 | -67.7062
0.13333 | -68.0301
0.13704 | -68.1191
0.14074 | -69.0437
0.14444 | -69.242
0.14815 | -70.2987
0.15185 | -69.6657
0.15556 | -70.7967
0.15926 | -70.4017
0.16296 | -70.2686
0.16667 | -69.901
0.17037 | -69.6296
0.17407 | -70.9155
0.17778 | -70.7605
0.18148 | -71.112
0.18518 | -71.4766
0.18889 | -70.821
0.19259 | -71.2206
0.1963 | -71.5464
0.2 | -71.0905
0.2037 | -71.7245
0.20741 | -71.9035
0.21111 | -71.8464
0.21482 | -73.6721
0.21852 | -72.3031
0.22222 | -71.6554
0.22593 | -71.3396
0.22963 | -73.4331
0.23333 | -73.4866
0.23704 | -74.157
0.24074 | -74.3117
0.24444 | -73.4812
0.24815 | -73.2873
0.25185 | -73.3715
0.25556 | -73.4391
0.25926 | -74.1863
0.26296 | -74.6518
0.26667 | -74.1784
0.27037 | -75.4525
0.27407 | -75.6195
0.27778 | -74.999
0.28148 | -75.2536
0.28519 | -75.8534
0.28889 | -76.1673
0.29259 | -77.3034
0.2963 | -76.9539
0.3 | -75.8517
0.3037 | -75.8868
0.30741 | -75.0103
0.31111 | -75.7869
0.31482 | -77.6317
0.31852 | -77.5798
0.32222 | -78.2165
0.32593 | -78.5563
0.32963 | -77.1296
0.33333 | -76.2249
0.33704 | -76.3001
0.34074 | -75.6791
0.34444 | -75.7392
0.34815 | -75.8242
0.35185 | -76.4475
0.35556 | -77.0314
0.35926 | -77.2254
0.36296 | -77.8835
0.36667 | -77.3695
0.37037 | -78.6241
0.37407 | -78.2136
0.37778 | -77.6131
0.38148 | -77.4663
0.38518 | -78.5884
0.38889 | -79.3207
0.39259 | -79.4181
0.3963 | -79.4635
0.4 | -79.4769
0.4037 | -80.3262
0.40741 | -79.7222
0.41111 | -78.6792
0.41481 | -78.846
0.41852 | -78.6837
0.42222 | -79.1698
0.42593 | -80.1663
0.42963 | -80.6892
0.43333 | -80.8222
0.43704 | -80.3871
0.44074 | -79.8525
0.44444 | -80.1066
0.44815 | -80.064
0.45185 | -79.9478
0.45556 | -80.2998
0.45926 | -80.5339
0.46296 | -79.1055
0.46667 | -80.0402
0.47037 | -80.3327
0.47407 | -81.4857
0.47778 | -81.5714
0.48148 | -80.7393
0.48518 | -81.1643
0.48889 | -80.7066
0.49259 | -80.8488
0.4963 | -80.267
0.5 | -80.8248
0.5037 | -81.9058
0.50741 | -81.5159
0.51111 | -81.0859
0.51482 | -81.8551
0.51852 | -81.0819
0.52222 | -80.4496
0.52593 | -80.6097
0.52963 | -80.1504
0.53333 | -80.6741
0.53704 | -81.2499
0.54074 | -82.0724
0.54444 | -81.9412
0.54815 | -81.3558
0.55185 | -81.6218
0.55556 | -82.1683
0.55926 | -82.0005
0.56296 | -81.7052
0.56667 | -81.516
0.57037 | -81.3361
0.57407 | -81.4951
0.57778 | -82.3379
0.58148 | -83.2351
0.58518 | -83.4269
0.58889 | -83.1701
0.59259 | -83.1714
0.5963 | -83.2249
0.6 | -82.5843
0.6037 | -81.9315
0.60741 | -82.0219
0.61111 | -82.4075
0.61482 | -83.0584
0.61852 | -83.5763
0.62222 | -83.5461
0.62593 | -83.2102
0.62963 | -82.7745
0.63333 | -82.5113
0.63704 | -82.5427
0.64074 | -82.7412
0.64444 | -83.1719
0.64815 | -83.4449
0.65185 | -83.294
0.65556 | -82.5986
0.65926 | -81.9727
0.66296 | -82.4474
0.66667 | -83.5467
0.67037 | -84.4202
0.67407 | -84.6657
0.67778 | -84.4254
0.68148 | -84.1124
0.68519 | -83.8727
0.68889 | -83.7694
0.69259 | -83.7163
0.6963 | -83.5266
0.7 | -83.2642
0.7037 | -83.4032
0.70741 | -84.2017
0.71111 | -85.1978
0.71481 | -85.3643
0.71852 | -84.6209
0.72222 | -84.1609
0.72593 | -84.2148
0.72963 | -84.6962
0.73333 | -84.8659
0.73704 | -84.3798
0.74074 | -83.7141
0.74444 | -83.4588
0.74815 | -83.8249
0.75185 | -84.5081
0.75556 | -84.9287
0.75926 | -84.7643
0.76296 | -84.6522
0.76667 | -84.8054
0.77037 | -85.059
0.77407 | -85.5818
0.77778 | -85.7824
0.78148 | -85.3584
0.78519 | -85.275
0.78889 | -85.551
0.79259 | -85.4606
0.7963 | -85.1433
0.8 | -85.2905
0.8037 | -85.6676
0.80741 | -85.8274
0.81111 | -85.8325
0.81481 | -86.1413
0.81852 | -85.9051
0.82222 | -85.5021
0.82593 | -85.4384
0.82963 | -85.5961
0.83333 | -85.969
0.83704 | -85.8152
0.84074 | -85.6613
0.84444 | -85.8951
0.84815 | -86.3344
0.85185 | -86.4615
0.85556 | -85.9966
0.85926 | -85.7453
0.86296 | -86.1139
0.86667 | -86.2493
0.87037 | -85.9768
0.87407 | -85.7649
0.87778 | -85.5778
0.88148 | -85.5569
0.88518 | -85.5407
0.88889 | -85.8128
0.89259 | -86.4747
0.8963 | -87.2406
0.9 | -87.7688
0.9037 | -87.4253
0.90741 | -86.6946
0.91111 | -86.2848
0.91482 | -86.3492
0.91852 | -86.6834
0.92222 | -86.8979
0.92593 | -86.8197
0.92963 | -86.93
0.93333 | -87.3464
0.93704 | -87.2894
0.94074 | -86.9844
0.94444 | -87.0193
0.94815 | -87.1715
0.95185 | -87.2917
0.95556 | -87.4444
0.95926 | -87.5824
0.96296 | -87.3994
0.96667 | -87.0455
0.97037 | -86.9913
0.97407 | -86.9945
0.97778 | -87.0363
0.98148 | -86.9916
0.98518 | -86.9607
0.98889 | -87.1053
0.99259 | -87.5776
0.9963 | -87.6586
1 | -87.2435
1.0037 | -87.353
1.0074 | -87.4126
1.0111 | -87.2971
1.0148 | -87.5231
1.0185 | -87.7645
1.0222 | -87.6642
1.0259 | -87.7717
1.0296 | -88.1429
1.0333 | -88.5058
1.037 | -88.4501
1.0407 | -88.2404
1.0444 | -88.1077
1.0481 | -88.2606
1.0519 | -88.4935
1.0556 | -88.4022
1.0593 | -88.214
1.063 | -87.7812
1.0667 | -87.4722
1.0704 | -87.3664
1.0741 | -87.4379
1.0778 | -87.5906
1.0815 | -87.6993
1.0852 | -87.7161
1.0889 | -87.7903
1.0926 | -88.0286
1.0963 | -88.2431
1.1 | -88.154
1.1037 | -87.9636
1.1074 | -88.1002
1.1111 | -88.5415
1.1148 | -88.9195
1.1185 | -89.0488
1.1222 | -88.9214
1.1259 | -88.5864
1.1296 | -88.1756
1.1333 | -87.8114
1.137 | -87.5529
1.1407 | -87.5005
1.1444 | -87.7136
1.1481 | -88.0357
1.1519 | -88.247
1.1556 | -88.2875
1.1593 | -88.2704
1.163 | -88.3899
1.1667 | -88.619
1.1704 | -88.7708
1.1741 | -88.8181
1.1778 | -88.8491
1.1815 | -88.853
1.1852 | -88.803
1.1889 | -88.8629
1.1926 | -89.1494
1.1963 | -89.4787
1.2 | -89.6086
1.2037 | -89.5409
1.2074 | -89.4472
1.2111 | -89.4135
1.2148 | -89.38
1.2185 | -89.2278
1.2222 | -88.901
1.2259 | -88.6454
1.2296 | -88.7598
1.2333 | -89.1616
1.237 | -89.535
1.2407 | -89.7254
1.2444 | -89.7384
1.2481 | -89.6835
1.2519 | -89.5658
1.2556 | -89.3636
1.2593 | -89.3186
1.263 | -89.5496
1.2667 | -89.7132
1.2704 | -89.575
1.2741 | -89.4844
1.2778 | -89.6487
1.2815 | -89.8342
1.2852 | -89.9492
1.2889 | -90.0582
1.2926 | -90.0498
1.2963 | -89.7519
1.3 | -89.3149
1.3037 | -89.1246
1.3074 | -89.2703
1.3111 | -89.4903
1.3148 | -89.5183
1.3185 | -89.3092
1.3222 | -89.0772
1.3259 | -89.1235
1.3296 | -89.518
1.3333 | -89.9777
1.337 | -90.1206
1.3407 | -89.9674
1.3444 | -89.8663
1.3481 | -89.8582
1.3519 | -89.7859
1.3556 | -89.7615
1.3593 | -89.8692
1.363 | -90.0478
1.3667 | -90.185
1.3704 | -90.124
1.3741 | -89.9486
1.3778 | -89.8776
1.3815 | -89.9053
1.3852 | -89.9769
1.3889 | -90.1655
1.3926 | -90.4397
1.3963 | -90.5794
1.4 | -90.5539
1.4037 | -90.5495
1.4074 | -90.6616
1.4111 | -90.7941
1.4148 | -90.7811
1.4185 | -90.8032
1.4222 | -91.0483
1.4259 | -91.2338
1.4296 | -91.0333
1.4333 | -90.678
1.437 | -90.4417
1.4407 | -90.2713
1.4444 | -90.1759
1.4481 | -90.3169
1.4519 | -90.6056
1.4556 | -90.8391
1.4593 | -91.0281
1.463 | -91.2175
1.4667 | -91.2997
1.4704 | -91.1637
1.4741 | -90.8583
1.4778 | -90.6137
1.4815 | -90.6458
1.4852 | -90.8986
1.4889 | -91.1303
1.4926 | -91.2263
1.4963 | -91.2738
1.5 | -91.3482
1.5037 | -91.4226
1.5074 | -91.354
1.5111 | -91.0426
1.5148 | -90.678
1.5185 | -90.5139
1.5222 | -90.5017
1.5259 | -90.4434
1.5296 | -90.4101
1.5333 | -90.5388
1.537 | -90.8072
1.5407 | -91.1249
1.5444 | -91.2995
1.5481 | -91.2853
1.5519 | -91.312
1.5556 | -91.4354
1.5593 | -91.5953
1.563 | -91.7447
1.5667 | -91.8007
1.5704 | -91.7376
1.5741 | -91.7012
1.5778 | -91.7991
1.5815 | -91.8833
1.5852 | -91.9324
1.5889 | -92.101
1.5926 | -92.1158
1.5963 | -91.6893
1.6 | -91.2921
1.6037 | -91.0833
1.6074 | -91.0181
1.6111 | -91.0298
1.6148 | -91.0816
1.6185 | -91.2219
1.6222 | -91.458
1.6259 | -91.7121
1.6296 | -91.9336
1.6333 | -92.1092
1.637 | -92.1803
1.6407 | -92.0625
1.6444 | -91.7829
1.6481 | -91.4715
1.6519 | -91.2412
1.6556 | -91.1468
1.6593 | -91.18
1.663 | -91.2567
1.6667 | -91.3073
1.6704 | -91.2596
1.6741 | -91.121
1.6778 | -90.9818
1.6815 | -90.9683
1.6852 | -91.1554
1.6889 | -91.5114
1.6926 | -91.8946
1.6963 | -92.124
1.7 | -92.105
1.7037 | -91.9821
1.7074 | -91.9412
1.7111 | -92.0467
1.7148 | -92.1772
1.7185 | -92.195
1.7222 | -92.1128
1.7259 | -91.99
1.7296 | -91.8528
1.7333 | -91.7287
1.737 | -91.7039
1.7407 | -91.8047
1.7444 | -91.9477
1.7481 | -92.0817
1.7519 | -92.2085
1.7556 | -92.3271
1.7593 | -92.4496
1.763 | -92.593
1.7667 | -92.7487
1.7704 | -92.8693
1.7741 | -92.9515
1.7778 | -93.0031
1.7815 | -92.9831
1.7852 | -92.8875
1.7889 | -92.7546
1.7926 | -92.5845
1.7963 | -92.3756
1.8 | -92.1666
1.8037 | -92.0167
1.8074 | -91.9617
1.8111 | -92.0153
1.8148 | -92.1708
1.8185 | -92.3426
1.8222 | -92.4371
1.8259 | -92.466
1.8296 | -92.5409
1.8333 | -92.7212
1.837 | -92.9168
1.8407 | -93.0013
1.8444 | -93.0074
1.8481 | -93.0651
1.8519 | -93.2128
1.8556 | -93.2758
1.8593 | -93.0994
1.863 | -92.7895
1.8667 | -92.5669
1.8704 | -92.5369
1.8741 | -92.6784
1.8778 | -92.9027
1.8815 | -93.1047
1.8852 | -93.178
1.8889 | -93.0909
1.8926 | -92.9223
1.8963 | -92.8128
1.9 | -92.883
1.9037 | -93.1209
1.9074 | -93.3899
1.9111 | -93.5278
1.9148 | -93.4721
1.9185 | -93.361
1.9222 | -93.3615
1.9259 | -93.4981
1.9296 | -93.658
1.9333 | -93.686
1.937 | -93.5723
1.9407 | -93.4032
1.9444 | -93.2593
1.9481 | -93.1725
1.9519 | -93.169
1.9556 | -93.3085
1.9593 | -93.6019
1.963 | -93.9305
1.9667 | -94.0707
1.9704 | -93.8818
1.9741 | -93.4931
1.9778 | -93.1298
1.9815 | -92.9439
1.9852 | -92.9557
1.9889 | -93.0936
1.9926 | -93.2783
1.9963 | -93.4862
2 | -93.7017
2.0037 | -93.8668
2.0074 | -93.9009
2.0111 | -93.8006
2.0148 | -93.6315
2.0185 | -93.4553
2.0222 | -93.32
2.0259 | -93.2531
2.0296 | -93.278
2.0333 | -93.4003
2.037 | -93.5419
2.0407 | -93.5482
2.0444 | -93.3574
2.0481 | -93.1216
2.0519 | -93.032
2.0556 | -93.1191
2.0593 | -93.2903
2.063 | -93.4448
2.0667 | -93.5756
2.0704 | -93.7005
2.0741 | -93.7801
2.0778 | -93.7703
2.0815 | -93.7192
2.0852 | -93.7378
2.0889 | -93.9086
2.0926 | -94.2089
2.0963 | -94.5096
2.1 | -94.6799
2.1037 | -94.7201
2.1074 | -94.7033
2.1111 | -94.6423
2.1148 | -94.5528
2.1185 | -94.4542
2.1222 | -94.3339
2.1259 | -94.1787
2.1296 | -93.9534
2.1333 | -93.6558
2.137 | -93.386
2.1407 | -93.1832
2.1444 | -93.0485
2.1481 | -92.998
2.1519 | -93.028
2.1556 | -93.1459
2.1593 | -93.3331
2.163 | -93.5221
2.1667 | -93.6662
2.1704 | -93.7571
2.1741 | -93.823
2.1778 | -93.9076
2.1815 | -94.018
2.1852 | -94.1257
2.1889 | -94.2027
2.1926 | -94.2515
2.1963 | -94.3018
2.2 | -94.3805
2.2037 | -94.4824
2.2074 | -94.5923
2.2111 | -94.7066
2.2148 | -94.7979
2.2185 | -94.8249
2.2222 | -94.7696
2.2259 | -94.6324
2.2296 | -94.4181
2.2333 | -94.1649
2.237 | -93.9052
2.2407 | -93.7028
2.2444 | -93.6188
2.2481 | -93.6634
2.2519 | -93.8071
2.2556 | -93.9993
2.2593 | -94.1602
2.263 | -94.2554
2.2667 | -94.3006
2.2704 | -94.2866
2.2741 | -94.2127
2.2778 | -94.0918
2.2815 | -93.9266
2.2852 | -93.7518
2.2889 | -93.6239
2.2926 | -93.5922
2.2963 | -93.6865
2.3 | -93.9118
2.3037 | -94.2198
2.3074 | -94.5222
2.3111 | -94.7567
2.3148 | -94.8554
2.3185 | -94.8243
2.3222 | -94.7501
2.3259 | -94.6815
2.3296 | -94.6489
2.3333 | -94.6551
2.337 | -94.6953
2.3407 | -94.7722
2.3444 | -94.8711
2.3481 | -94.9797
2.3519 | -95.0729
2.3556 | -95.1409
2.3593 | -95.1697
2.363 | -95.1185
2.3667 | -95.0103
2.3704 | -94.9065
2.3741 | -94.8747
2.3778 | -94.9243
2.3815 | -94.9987
2.3852 | -95.021
2.3889 | -94.9304
2.3926 | -94.753
2.3963 | -94.557
2.4 | -94.4002
2.4037 | -94.3164
2.4074 | -94.2948
2.4111 | -94.329
2.4148 | -94.4262
2.4185 | -94.5819
2.4222 | -94.7735
2.4259 | -94.9588
2.4296 | -95.0851
2.4333 | -95.1046
2.437 | -95.0249
2.4407 | -94.9141
2.4444 | -94.862
2.4481 | -94.9159
2.4519 | -95.0572
2.4556 | -95.2252
2.4593 | -95.3682
2.463 | -95.4612
2.4667 | -95.4889
2.4704 | -95.4354
2.4741 | -95.3057
2.4778 | -95.099
2.4815 | -94.8452
2.4852 | -94.6234
2.4889 | -94.513
2.4926 | -94.5432
2.4963 | -94.6744
2.5 | -94.8266
2.5037 | -94.9246
2.5074 | -94.972
2.5111 | -95.0071
2.5148 | -95.0614
2.5185 | -95.148
2.5222 | -95.2371
2.5259 | -95.3056
2.5296 | -95.3412
2.5333 | -95.35
2.537 | -95.3244
2.5407 | -95.2391
2.5444 | -95.0973
2.5481 | -94.9196
2.5519 | -94.7464
2.5556 | -94.6079
2.5593 | -94.5458
2.563 | -94.58
2.5667 | -94.6977
2.5704 | -94.8644
2.5741 | -95.0107
2.5778 | -95.0741
2.5815 | -95.0421
2.5852 | -94.9558
2.5889 | -94.8604
2.5926 | -94.7977
2.5963 | -94.8033
2.6 | -94.897
2.6037 | -95.0675
2.6074 | -95.2456
2.6111 | -95.3589
2.6148 | -95.374
2.6185 | -95.3302
2.6222 | -95.3122
2.6259 | -95.3548
2.6296 | -95.4554
2.6333 | -95.5618
2.637 | -95.5956
2.6407 | -95.5109
2.6444 | -95.3447
2.6481 | -95.1762
2.6519 | -95.0759
2.6556 | -95.0643
2.6593 | -95.1144
2.663 | -95.1581
2.6667 | -95.195
2.6704 | -95.3034
2.6741 | -95.4321
2.6778 | -95.59
2.6815 | -95.7362
2.6852 | -95.823
2.6889 | -95.8124
2.6926 | -95.7192
2.6963 | -95.5934
2.7 | -95.4869
2.7037 | -95.4376
2.7074 | -95.4673
2.7111 | -95.5595
2.7148 | -95.6821
2.7185 | -95.8066
2.7222 | -95.9123
2.7259 | -95.9958
2.7296 | -96.0558
2.7333 | -96.0761
2.737 | -96.0394
2.7407 | -95.9574
2.7444 | -95.8504
2.7481 | -95.7372
2.7519 | -95.628
2.7556 | -95.525
2.7593 | -95.4064
2.763 | -95.2493
2.7667 | -95.0519
2.7704 | -94.8439
2.7741 | -94.6679
2.7778 | -94.5692
2.7815 | -94.5771
2.7852 | -94.7072
2.7889 | -94.9392
2.7926 | -95.2519
2.7963 | -95.6076
2.8 | -95.9468
2.8037 | -96.2285
2.8074 | -96.4381
2.8111 | -96.5621
2.8148 | -96.5936
2.8185 | -96.5509
2.8222 | -96.4771
2.8259 | -96.4041
2.8296 | -96.3554
2.8333 | -96.3154
2.837 | -96.2493
2.8407 | -96.1433
2.8444 | -96.0105
2.8481 | -95.8922
2.8519 | -95.8066
2.8556 | -95.7694
2.8593 | -95.7802
2.863 | -95.8196
2.8667 | -95.8852
2.8704 | -95.9783
2.8741 | -96.1018
2.8778 | -96.2353
2.8815 | -96.3624
2.8852 | -96.4663
2.8889 | -96.5225
2.8926 | -96.5392
2.8963 | -96.5196
2.9 | -96.4911
2.9037 | -96.4623
2.9074 | -96.4384
2.9111 | -96.4335
2.9148 | -96.4393
2.9185 | -96.4456
2.9222 | -96.439
2.9259 | -96.4015
2.9296 | -96.3295
2.9333 | -96.2446
2.937 | -96.184
2.9407 | -96.1452
2.9444 | -96.1336
2.9481 | -96.1122
2.9519 | -96.0551
2.9556 | -95.9545
2.9593 | -95.7995
2.963 | -95.6231
2.9667 | -95.4689
2.9704 | -95.3793
2.9741 | -95.3719
2.9778 | -95.4316
2.9815 | -95.5029
2.9852 | -95.5363
2.9889 | -95.5368
2.9926 | -95.5318
2.9963 | -95.5722
3 | -95.6693
3.0037 | -95.8144
3.0074 | -95.9965
3.0111 | -96.1856
3.0148 | -96.3658
3.0185 | -96.5298
3.0222 | -96.6625
3.0259 | -96.745
3.0296 | -96.7949
3.0333 | -96.8296
3.037 | -96.8563
3.0407 | -96.8712
3.0444 | -96.8685
3.0481 | -96.8357
3.0519 | -96.7699
3.0556 | -96.6994
3.0593 | -96.665
3.063 | -96.7052
3.0667 | -96.8299
3.0704 | -97.0236
3.0741 | -97.2332
3.0778 | -97.3937
3.0815 | -97.4619
3.0852 | -97.4409
3.0889 | -97.3794
3.0926 | -97.3149
3.0963 | -97.2673
3.1 | -97.2269
3.1037 | -97.1773
3.1074 | -97.1186
3.1111 | -97.0417
3.1148 | -96.9525
3.1185 | -96.8559
3.1222 | -96.7677
3.1259 | -96.7105
3.1296 | -96.6903
3.1333 | -96.7053
3.137 | -96.7365
3.1407 | -96.7785
3.1444 | -96.8251
3.1481 | -96.8834
3.1519 | -96.9337
3.1556 | -96.9474
3.1593 | -96.9305
3.163 | -96.8871
3.1667 | -96.842
3.1704 | -96.8178
3.1741 | -96.8171
3.1778 | -96.8485
3.1815 | -96.9016
3.1852 | -96.9557
3.1889 | -96.9807
3.1926 | -96.9865
3.1963 | -96.9766
3.2 | -96.9827
3.2037 | -97.0541
3.2074 | -97.1212
3.2111 | -97.1948
3.2148 | -97.2719
3.2185 | -97.3352
3.2222 | -97.3738
3.2259 | -97.3836
3.2296 | -97.3683
3.2333 | -97.3172
3.237 | -97.2584
3.2407 | -97.2114
3.2444 | -97.1754
3.2481 | -97.1695
3.2519 | -97.2
3.2556 | -97.26
3.2593 | -97.3363
3.263 | -97.4241
3.2667 | -97.5064
3.2704 | -97.5706
3.2741 | -97.6243
3.2778 | -97.6736
3.2815 | -97.7227
3.2852 | -97.7519
3.2889 | -97.762
3.2926 | -97.7282
3.2963 | -97.6515
3.3 | -97.5641
3.3037 | -97.4815
3.3074 | -97.4226
3.3111 | -97.4056
3.3148 | -97.4246
3.3185 | -97.4735
3.3222 | -97.5526
3.3259 | -97.6447
3.3296 | -97.7223
3.3333 | -97.7693
3.337 | -97.769
3.3407 | -97.7082
3.3444 | -97.5856
3.3481 | -97.4329
3.3519 | -97.286
3.3556 | -97.1674
3.3593 | -97.1015
3.363 | -97.0897
3.3667 | -97.1289
3.3704 | -97.2065
3.3741 | -97.2883
3.3778 | -97.3647
3.3815 | -97.428
3.3852 | -97.4787
3.3889 | -97.5145
3.3926 | -97.5417
3.3963 | -97.5677
3.4 | -97.5868
3.4037 | -97.5844
3.4074 | -97.548
3.4111 | -97.492
3.4148 | -97.4355
3.4185 | -97.3879
3.4222 | -97.3576
3.4259 | -97.3565
3.4296 | -97.3937
3.4333 | -97.4574
3.437 | -97.5356
3.4407 | -97.6222
3.4444 | -97.6983
3.4481 | -97.7309
3.4519 | -97.7084
3.4556 | -97.6412
3.4593 | -97.5524
3.463 | -97.485
3.4667 | -97.4611
3.4704 | -97.49
3.4741 | -97.5738
3.4778 | -97.6879
3.4815 | -97.8241
3.4852 | -97.9598
3.4889 | -98.0888
3.4926 | -98.204
3.4963 | -98.2892
3.5 | -98.3482
3.5037 | -98.3796
3.5074 | -98.3785
3.5111 | -98.3547
3.5148 | -98.3083
3.5185 | -98.2322
3.5222 | -98.1323
3.5259 | -98.0444
3.5296 | -97.9909
3.5333 | -97.982
3.537 | -98.0055
3.5407 | -98.054
3.5444 | -98.0942
3.5481 | -98.1121
3.5519 | -98.0953
3.5556 | -98.0493
3.5593 | -97.9897
3.563 | -97.9189
3.5667 | -97.8444
3.5704 | -97.7831
3.5741 | -97.7212
3.5778 | -97.6531
3.5815 | -97.572
3.5852 | -97.4839
3.5889 | -97.4097
3.5926 | -97.3633
3.5963 | -97.3537
3.6 | -97.3763
3.6037 | -97.4181
3.6074 | -97.4797
3.6111 | -97.5579
3.6148 | -97.661
3.6185 | -97.7976
3.6222 | -97.9628
3.6259 | -98.1495
3.6296 | -98.349
3.6333 | -98.5322
3.637 | -98.6787
3.6407 | -98.7858
3.6444 | -98.8656
3.6481 | -98.9245
3.6519 | -98.9625
3.6556 | -98.9732
3.6593 | -98.947
3.663 | -98.8614
3.6667 | -98.7433
3.6704 | -98.6066
3.6741 | -98.468
3.6778 | -98.3529
3.6815 | -98.2763
3.6852 | -98.2412
3.6889 | -98.2328
3.6926 | -98.2466
3.6963 | -98.2791
3.7 | -98.3111
3.7037 | -98.3442
3.7074 | -98.3862
3.7111 | -98.4399
3.7148 | -98.481
3.7185 | -98.5166
3.7222 | -98.5354
3.7259 | -98.5362
3.7296 | -98.514
3.7333 | -98.4659
3.737 | -98.3906
3.7407 | -98.3319
3.7444 | -98.2793
3.7481 | -98.2528
3.7519 | -98.2566
3.7556 | -98.2867
3.7593 | -98.3589
3.763 | -98.4579
3.7667 | -98.5473
3.7704 | -98.612
3.7741 | -98.6616
3.7778 | -98.6848
3.7815 | -98.6839
3.7852 | -98.6805
3.7889 | -98.6806
3.7926 | -98.6779
3.7963 | -98.6714
3.8 | -98.6553
3.8037 | -98.6334
3.8074 | -98.6092
3.8111 | -98.574
3.8148 | -98.5394
3.8185 | -98.5006
3.8222 | -98.4488
3.8259 | -98.4028
3.8296 | -98.3685
3.8333 | -98.3409
3.837 | -98.317
3.8407 | -98.3057
3.8444 | -98.3163
3.8481 | -98.3433
3.8519 | -98.3898
3.8556 | -98.4478
3.8593 | -98.5104
3.863 | -98.5688
3.8667 | -98.6124
3.8704 | -98.6344
3.8741 | -98.6489
3.8778 | -98.66
3.8815 | -98.6856
3.8852 | -98.7126
3.8889 | -98.7299
3.8926 | -98.7365
3.8963 | -98.7318
3.9 | -98.7209
3.9037 | -98.6937
3.9074 | -98.6518
3.9111 | -98.6107
3.9148 | -98.579
3.9185 | -98.5726
3.9222 | -98.5837
3.9259 | -98.6164
3.9296 | -98.6648
3.9333 | -98.7199
3.937 | -98.769
3.9407 | -98.802
3.9444 | -98.8194
3.9481 | -98.8261
3.9519 | -98.8187
3.9556 | -98.8135
3.9593 | -98.8045
3.963 | -98.8039
3.9667 | -98.8303
3.9704 | -98.8654
3.9741 | -98.9057
3.9778 | -98.953
3.9815 | -99.0097
3.9852 | -99.0927
3.9889 | -99.1855
3.9926 | -99.2835
3.9963 | -99.3948
4 | -99.4912
4.0037 | -99.544
4.0074 | -99.5359
4.0111 | -99.485
4.0148 | -99.4034
4.0185 | -99.3009
4.0222 | -99.2088
4.0259 | -99.1331
4.0296 | -99.0839
4.0333 | -99.0654
4.037 | -99.0539
4.0407 | -99.039
4.0444 | -99.0122
4.0481 | -98.9795
4.0519 | -98.9442
4.0556 | -98.9025
4.0593 | -98.8594
4.063 | -98.828
4.0667 | -98.8127
4.0704 | -98.8167
4.0741 | -98.8507
4.0778 | -98.9158
4.0815 | -99.0105
4.0852 | -99.123
4.0889 | -99.2447
4.0926 | -99.3515
4.0963 | -99.4414
4.1 | -99.5148
4.1037 | -99.5543
4.1074 | -99.5585
4.1111 | -99.5175
4.1148 | -99.4602
4.1185 | -99.3964
4.1222 | -99.3321
4.1259 | -99.2902
4.1296 | -99.2771
4.1333 | -99.2915
4.137 | -99.3161
4.1407 | -99.3498
4.1444 | -99.3814
4.1481 | -99.4053
4.1519 | -99.4063
4.1556 | -99.3938
4.1593 | -99.3692
4.163 | -99.3357
4.1667 | -99.315
4.1704 | -99.3018
4.1741 | -99.3063
4.1778 | -99.3207
4.1815 | -99.3373
4.1852 | -99.3602
4.1889 | -99.3766
4.1926 | -99.3909
4.1963 | -99.4104
4.2 | -99.4358
4.2037 | -99.472
4.2074 | -99.5102
4.2111 | -99.5387
4.2148 | -99.5616
4.2185 | -99.5769
4.2222 | -99.5884
4.2259 | -99.6136
4.2296 | -99.6313
4.2333 | -99.6454
4.237 | -99.6453
4.2407 | -99.6209
4.2444 | -99.5794
4.2481 | -99.5165
4.2519 | -99.4575
4.2556 | -99.4052
4.2593 | -99.3534
4.263 | -99.317
4.2667 | -99.2849
4.2704 | -99.248
4.2741 | -99.1899
4.2778 | -99.136
4.2815 | -99.0789
4.2852 | -99.0326
4.2889 | -99.0034
4.2926 | -98.9909
4.2963 | -99.0121
4.3 | -99.061
4.3037 | -99.1312
4.3074 | -99.2225
4.3111 | -99.3218
4.3148 | -99.4192
4.3185 | -99.4955
4.3222 | -99.5486
4.3259 | -99.5663
4.3296 | -99.542
4.3333 | -99.4863
4.337 | -99.4115
4.3407 | -99.3376
4.3444 | -99.2768
4.3481 | -99.2386
4.3519 | -99.2214
4.3556 | -99.2206
4.3593 | -99.234
4.363 | -99.2658
4.3667 | -99.3046
4.3704 | -99.3428
4.3741 | -99.378
4.3778 | -99.4019
4.3815 | -99.4239
4.3852 | -99.4322
4.3889 | -99.422
4.3926 | -99.4107
4.3963 | -99.4053
4.4 | -99.3915
4.4037 | -99.3647
4.4074 | -99.3275
4.4111 | -99.2786
4.4148 | -99.2303
4.4185 | -99.1949
4.4222 | -99.1831
4.4259 | -99.1985
4.4296 | -99.2477
4.4333 | -99.3208
4.437 | -99.4136
4.4407 | -99.5286
4.4444 | -99.6582
4.4481 | -99.7926
4.4519 | -99.9172
4.4556 | -100.0289
4.4593 | -100.1196
4.463 | -100.1757
4.4667 | -100.2034
4.4704 | -100.2129
4.4741 | -100.2117
4.4778 | -100.1939
4.4815 | -100.1695
4.4852 | -100.1469
4.4889 | -100.1236
4.4926 | -100.1097
4.4963 | -100.1191
4.5 | -100.1493
4.5037 | -100.2018
4.5074 | -100.2696
4.5111 | -100.337
4.5148 | -100.3969
4.5185 | -100.445
4.5222 | -100.4732
4.5259 | -100.486
4.5296 | -100.482
4.5333 | -100.4611
4.537 | -100.435
4.5407 | -100.4104
4.5444 | -100.3889
4.5481 | -100.3821
4.5519 | -100.3817
4.5556 | -100.4023
4.5593 | -100.4513
4.563 | -100.5101
4.5667 | -100.5735
4.5704 | -100.6376
4.5741 | -100.6934
4.5778 | -100.7234
4.5815 | -100.718
4.5852 | -100.671
4.5889 | -100.5945
4.5926 | -100.4936
4.5963 | -100.3718
4.6 | -100.2405
4.6037 | -100.1224
4.6074 | -100.0308
4.6111 | -99.9678
4.6148 | -99.9282
4.6185 | -99.9045
4.6222 | -99.8912
4.6259 | -99.8744
4.6296 | -99.8521
4.6333 | -99.8257
4.637 | -99.7924
4.6407 | -99.7564
4.6444 | -99.7103
4.6481 | -99.6706
4.6519 | -99.6427
4.6556 | -99.6325
4.6593 | -99.6397
4.663 | -99.6704
4.6667 | -99.7312
4.6704 | -99.8084
4.6741 | -99.8834
4.6778 | -99.9499
4.6815 | -100.0049
4.6852 | -100.0536
4.6889 | -100.0932
4.6926 | -100.1213
4.6963 | -100.1401
4.7 | -100.155
4.7037 | -100.1604
4.7074 | -100.1645
4.7111 | -100.1625
4.7148 | -100.1594
4.7185 | -100.1518
4.7222 | -100.1421
4.7259 | -100.1472
4.7296 | -100.1676
4.7333 | -100.2047
4.737 | -100.2578
4.7407 | -100.326
4.7444 | -100.4031
4.7481 | -100.4878
4.7519 | -100.5692
4.7556 | -100.6444
4.7593 | -100.7048
4.763 | -100.7516
4.7667 | -100.7741
4.7704 | -100.7669
4.7741 | -100.7219
4.7778 | -100.6394
4.7815 | -100.545
4.7852 | -100.4556
4.7889 | -100.364
4.7926 | -100.2811
4.7963 | -100.2065
4.8 | -100.1481
4.8037 | -100.103
4.8074 | -100.1164
4.8111 | -100.1021
4.8148 | -100.1089
4.8185 | -100.1352
4.8222 | -100.1684
4.8259 | -100.2151
4.8296 | -100.2652
4.8333 | -100.3207
4.837 | -100.3802
4.8407 | -100.4409
4.8444 | -100.5031
4.8481 | -100.577
4.8519 | -100.6483
4.8556 | -100.7047
4.8593 | -100.7516
4.863 | -100.7925
4.8667 | -100.8213
4.8704 | -100.8361
4.8741 | -100.8305
4.8778 | -100.8155
4.8815 | -100.7869
4.8852 | -100.7488
4.8889 | -100.7189
4.8926 | -100.6771
4.8963 | -100.6291
4.9 | -100.5906
4.9037 | -100.563
4.9074 | -100.5475
4.9111 | -100.5519
4.9148 | -100.5674
4.9185 | -100.5968
4.9222 | -100.6302
4.9259 | -100.6662
4.9296 | -100.7081
4.9333 | -100.7495
4.937 | -100.7823
4.9407 | -100.8112
4.9444 | -100.8313
4.9481 | -100.8428
4.9519 | -100.8512
4.9556 | -100.8547
4.9593 | -100.853
4.963 | -100.8555
4.9667 | -100.8617
4.9704 | -100.879
4.9741 | -100.9044
4.9778 | -100.9279
4.9815 | -100.9517
4.9852 | -100.9674
4.9889 | -100.9753
4.9926 | -100.9865
4.9963 | -100.9901
5 | -100.9897
5.0037 | -100.9809
5.0074 | -100.9631
5.0111 | -100.9387
5.0148 | -100.9045
5.0185 | -100.8558
5.0222 | -100.8108
5.0259 | -100.7741
5.0296 | -100.7385
5.0333 | -100.7004
5.037 | -100.6662
5.0407 | -100.6319
5.0444 | -100.5873
5.0481 | -100.5348
5.0519 | -100.4803
5.0556 | -100.4305
5.0593 | -100.3805
5.063 | -100.3431
5.0667 | -100.3083
5.0704 | -100.2829
5.0741 | -100.2781
5.0778 | -100.2888
5.0815 | -100.3153
5.0852 | -100.3582
5.0889 | -100.4213
5.0926 | -100.5024
5.0963 | -100.5986
5.1 | -100.6872
5.1037 | -100.7766
5.1074 | -100.8702
5.1111 | -100.9549
5.1148 | -101.0306
5.1185 | -101.1007
5.1222 | -101.1629
5.1259 | -101.2138
5.1296 | -101.2488
5.1333 | -101.2695
5.137 | -101.2792
5.1407 | -101.276
5.1444 | -101.26
5.1481 | -101.2173
5.1519 | -101.1542
5.1556 | -101.0852
5.1593 | -101.0169
5.163 | -100.9486
5.1667 | -100.8794
5.1704 | -100.8192
5.1741 | -100.7761
5.1778 | -100.7451
5.1815 | -100.7182
5.1852 | -100.6984
5.1889 | -100.6864
5.1926 | -100.671
5.1963 | -100.6495
5.2 | -100.6276
5.2037 | -100.6124
5.2074 | -100.6013
5.2111 | -100.592
5.2148 | -100.5855
5.2185 | -100.5885
5.2222 | -100.6095
5.2259 | -100.6495
5.2296 | -100.7007
5.2333 | -100.759
5.237 | -100.8195
5.2407 | -100.8776
5.2444 | -100.9345
5.2481 | -100.9752
5.2519 | -100.9987
5.2556 | -101.0009
5.2593 | -100.9799
5.263 | -100.947
5.2667 | -100.9017
5.2704 | -100.8502
5.2741 | -100.802
5.2778 | -100.7544
5.2815 | -100.6978
5.2852 | -100.6489
5.2889 | -100.6105
5.2926 | -100.5775
5.2963 | -100.5454
5.3 | -100.5226
5.3037 | -100.5147
5.3074 | -100.5208
5.3111 | -100.5411
5.3148 | -100.5718
5.3185 | -100.6102
5.3222 | -100.6364
5.3259 | -100.6515
5.3296 | -100.6543
5.3333 | -100.6557
5.337 | -100.6515
5.3407 | -100.6554
5.3444 | -100.6654
5.3481 | -100.6923
5.3519 | -100.7318
5.3556 | -100.778
5.3593 | -100.8419
5.363 | -100.9215
5.3667 | -101.0007
5.3704 | -101.0785
5.3741 | -101.1557
5.3778 | -101.2302
5.3815 | -101.29
5.3852 | -101.3364
5.3889 | -101.3635
5.3926 | -101.3615
5.3963 | -101.3415
5.4 | -101.3096
5.4037 | -101.276
5.4074 | -101.2383
5.4111 | -101.1931
5.4148 | -101.1507
5.4185 | -101.1196
5.4222 | -101.094
5.4259 | -101.0686
5.4296 | -101.0488
5.4333 | -101.0344
5.437 | -101.0337
5.4407 | -101.0438
5.4444 | -101.0554
5.4481 | -101.0764
5.4519 | -101.1035
5.4556 | -101.1409
5.4593 | -101.1797
5.463 | -101.2097
5.4667 | -101.2395
5.4704 | -101.2594
5.4741 | -101.2687
5.4778 | -101.2768
5.4815 | -101.2722
5.4852 | -101.2538
5.4889 | -101.2259
5.4926 | -101.196
5.4963 | -101.1714
5.5 | -101.1543
5.5037 | -101.1502
5.5074 | -101.1626
5.5111 | -101.1964
5.5148 | -101.2416
5.5185 | -101.2913
5.5222 | -101.3414
5.5259 | -101.3955
5.5296 | -101.4396
5.5333 | -101.4681
5.537 | -101.4813
5.5407 | -101.4681
5.5444 | -101.4385
5.5481 | -101.4031
5.5519 | -101.3531
5.5556 | -101.3037
5.5593 | -101.2567
5.563 | -101.2126
5.5667 | -101.178
5.5704 | -101.1573
5.5741 | -101.155
5.5778 | -101.1676
5.5815 | -101.1879
5.5852 | -101.2089
5.5889 | -101.2355
5.5926 | -101.2549
5.5963 | -101.2712
5.6 | -101.2911
5.6037 | -101.3119
5.6074 | -101.3322
5.6111 | -101.3504
5.6148 | -101.3688
5.6185 | -101.3847
5.6222 | -101.3972
5.6259 | -101.4044
5.6296 | -101.4082
5.6333 | -101.4109
5.637 | -101.4246
5.6407 | -101.4448
5.6444 | -101.4714
5.6481 | -101.4988
5.6519 | -101.5349
5.6556 | -101.573
5.6593 | -101.6025
5.663 | -101.6287
5.6667 | -101.6527
5.6704 | -101.6635
5.6741 | -101.658
5.6778 | -101.6468
5.6815 | -101.6183
5.6852 | -101.5755
5.6889 | -101.5291
5.6926 | -101.4745
5.6963 | -101.4153
5.7 | -101.3545
5.7037 | -101.2896
5.7074 | -101.2296
5.7111 | -101.1729
5.7148 | -101.1241
5.7185 | -101.08
5.7222 | -101.0462
5.7259 | -101.0219
5.7296 | -101.0147
5.7333 | -101.02
5.737 | -101.0292
5.7407 | -101.042
5.7444 | -101.0597
5.7481 | -101.0785
5.7519 | -101.096
5.7556 | -101.1146
5.7593 | -101.1291
5.763 | -101.137
5.7667 | -101.1403
5.7704 | -101.1467
5.7741 | -101.1642
5.7778 | -101.1965
5.7815 | -101.237
5.7852 | -101.294
5.7889 | -101.3667
5.7926 | -101.4429
5.7963 | -101.5234
5.8 | -101.5969
5.8037 | -101.6498
5.8074 | -101.6808
5.8111 | -101.6907
5.8148 | -101.6751
5.8185 | -101.6353
5.8222 | -101.5823
5.8259 | -101.5158
5.8296 | -101.4393
5.8333 | -101.3596
5.837 | -101.288
5.8407 | -101.2199
5.8444 | -101.1628
5.8481 | -101.1161
5.8519 | -101.0749
5.8556 | -101.0326
5.8593 | -100.9947
5.863 | -100.9637
5.8667 | -100.9364
5.8704 | -100.9143
5.8741 | -100.8991
5.8778 | -100.8806
5.8815 | -100.8681
5.8852 | -100.8568
5.8889 | -100.8423
5.8926 | -100.8353
5.8963 | -100.8327
5.9 | -100.8255
5.9037 | -100.8122
5.9074 | -100.8022
5.9111 | -100.7941
5.9148 | -100.796
5.9185 | -100.7985
5.9222 | -100.8002
5.9259 | -100.797
5.9296 | -100.7969
5.9333 | -100.7956
5.937 | -100.7963
5.9407 | -100.8029
5.9444 | -100.814
5.9481 | -100.83
5.9519 | -100.8502
5.9556 | -100.8804
5.9593 | -100.9087
5.963 | -100.9443
5.9667 | -100.9863
5.9704 | -101.0269
5.9741 | -101.0657
5.9778 | -101.1014
5.9815 | -101.1412
5.9852 | -101.1737
5.9889 | -101.2019
5.9926 | -101.2299
5.9963 | -101.2523
6 | -101.2679
6.0037 | -101.278
6.0074 | -101.2811
6.0111 | -101.2868
6.0148 | -101.2941
6.0185 | -101.3048
6.0222 | -101.3065
6.0259 | -101.3103
6.0296 | -101.3176
6.0333 | -101.3234
6.037 | -101.3298
6.0407 | -101.3337
6.0444 | -101.3408
6.0481 | -101.3476
6.0519 | -101.3586
6.0556 | -101.3711
6.0593 | -101.3927
6.063 | -101.4226
6.0667 | -101.4654
6.0704 | -101.5206
6.0741 | -101.5757
6.0778 | -101.6451
6.0815 | -101.7206
6.0852 | -101.7983
6.0889 | -101.8785
6.0926 | -101.9499
6.0963 | -102.0125
6.1 | -102.0685
6.1037 | -102.1197
6.1074 | -102.158
6.1111 | -102.1889
6.1148 | -102.2167
6.1185 | -102.2264
6.1222 | -102.2253
6.1259 | -102.2146
6.1296 | -102.1987
6.1333 | -102.1857
6.137 | -102.1695
6.1407 | -102.1517
6.1444 | -102.1437
6.1481 | -102.1462
6.1519 | -102.1613
6.1556 | -102.1828
6.1593 | -102.2097
6.163 | -102.2452
6.1667 | -102.2869
6.1704 | -102.3191
6.1741 | -102.3474
6.1778 | -102.3736
6.1815 | -102.3914
6.1852 | -102.396
6.1889 | -102.3862
6.1926 | -102.3639
6.1963 | -102.3329
6.2 | -102.299
6.2037 | -102.2599
6.2074 | -102.225
6.2111 | -102.1952
6.2148 | -102.1746
6.2185 | -102.1707
6.2222 | -102.1748
6.2259 | -102.1841
6.2296 | -102.2035
6.2333 | -102.2309
6.237 | -102.256
6.2407 | -102.2796
6.2444 | -102.2974
6.2481 | -102.3039
6.2519 | -102.3039
6.2556 | -102.2917
6.2593 | -102.2718
6.263 | -102.2434
6.2667 | -102.2088
6.2704 | -102.1718
6.2741 | -102.1281
6.2778 | -102.0845
6.2815 | -102.0444
6.2852 | -102.0096
6.2889 | -101.9761
6.2926 | -101.9483
6.2963 | -101.9267
6.3 | -101.9064
6.3037 | -101.8934
6.3074 | -101.8867
6.3111 | -101.8877
6.3148 | -101.8964
6.3185 | -101.9181
6.3222 | -101.9486
6.3259 | -101.9864
6.3296 | -102.0322
6.3333 | -102.0886
6.337 | -102.1526
6.3407 | -102.2266
6.3444 | -102.2969
6.3481 | -102.3638
6.3519 | -102.4236
6.3556 | -102.4795
6.3593 | -102.5271
6.363 | -102.5735
6.3667 | -102.6156
6.3704 | -102.645
6.3741 | -102.6559
6.3778 | -102.6566
6.3815 | -102.6545
6.3852 | -102.6473
6.3889 | -102.6299
6.3926 | -102.6089
6.3963 | -102.576
6.4 | -102.5394
6.4037 | -102.5068
6.4074 | -102.4772
6.4111 | -102.4592
6.4148 | -102.4387
6.4185 | -102.431
6.4222 | -102.4301
6.4259 | -102.4253
6.4296 | -102.4197
6.4333 | -102.4186
6.437 | -102.4127
6.4407 | -102.4096
6.4444 | -102.4041
6.4481 | -102.3934
6.4519 | -102.3803
6.4556 | -102.3697
6.4593 | -102.3581
6.463 | -102.351
6.4667 | -102.3534
6.4704 | -102.3625
6.4741 | -102.3699
6.4778 | -102.3791
6.4815 | -102.3958
6.4852 | -102.4228
6.4889 | -102.4533
6.4926 | -102.4914
6.4963 | -102.5286
6.5 | -102.5636
6.5037 | -102.6079
6.5074 | -102.6454
6.5111 | -102.6713
6.5148 | -102.6934
6.5185 | -102.701
6.5222 | -102.7007
6.5259 | -102.6986
6.5296 | -102.6897
6.5333 | -102.6716
6.537 | -102.647
6.5407 | -102.6209
6.5444 | -102.601
6.5481 | -102.5874
6.5519 | -102.5802
6.5556 | -102.5733
6.5593 | -102.5708
6.563 | -102.5711
6.5667 | -102.5704
6.5704 | -102.5667
6.5741 | -102.5651
6.5778 | -102.5623
6.5815 | -102.5562
6.5852 | -102.5472
6.5889 | -102.5391
6.5926 | -102.5324
6.5963 | -102.5324
6.6 | -102.5321
6.6037 | -102.5269
6.6074 | -102.5308
6.6111 | -102.5349
6.6148 | -102.5355
6.6185 | -102.5437
6.6222 | -102.559
6.6259 | -102.575
6.6296 | -102.5905
6.6333 | -102.6103
6.637 | -102.6334
6.6407 | -102.6489
6.6444 | -102.659
6.6481 | -102.6682
6.6519 | -102.6714
6.6556 | -102.6756
6.6593 | -102.6779
6.663 | -102.6751
6.6667 | -102.6739
6.6704 | -102.6737
6.6741 | -102.6669
6.6778 | -102.6602
6.6815 | -102.6575
6.6852 | -102.6608
6.6889 | -102.6683
6.6926 | -102.6805
6.6963 | -102.6949
6.7 | -102.712
6.7037 | -102.7345
6.7074 | -102.7538
6.7111 | -102.7669
6.7148 | -102.7803
6.7185 | -102.7866
6.7222 | -102.7857
6.7259 | -102.7846
6.7296 | -102.7816
6.7333 | -102.7697
6.737 | -102.7553
6.7407 | -102.7389
6.7444 | -102.7201
6.7481 | -102.6989
6.7519 | -102.6788
6.7556 | -102.6566
6.7593 | -102.6374
6.763 | -102.6282
6.7667 | -102.6188
6.7704 | -102.6112
6.7741 | -102.6147
6.7778 | -102.6258
6.7815 | -102.6443
6.7852 | -102.6668
6.7889 | -102.6939
6.7926 | -102.7228
6.7963 | -102.7558
6.8 | -102.7951
6.8037 | -102.8313
6.8074 | -102.873
6.8111 | -102.9237
6.8148 | -102.967
6.8185 | -103.0124
6.8222 | -103.059
6.8259 | -103.0984
6.8296 | -103.1334
6.8333 | -103.1618
6.837 | -103.1749
6.8407 | -103.1805
6.8444 | -103.1806
6.8481 | -103.1647
6.8519 | -103.1318
6.8556 | -103.0942
6.8593 | -103.0616
6.863 | -103.0214
6.8667 | -102.9735
6.8704 | -102.925
6.8741 | -102.8637
6.8778 | -102.7904
6.8815 | -102.7141
6.8852 | -102.6449
6.8889 | -102.5826
6.8926 | -102.5229
6.8963 | -102.4615
6.9 | -102.4169
6.9037 | -102.3888
6.9074 | -102.373
6.9111 | -102.3682
6.9148 | -102.3803
6.9185 | -102.4055
6.9222 | -102.4416
6.9259 | -102.4855
6.9296 | -102.5362
6.9333 | -102.5966
6.937 | -102.6609
6.9407 | -102.716
6.9444 | -102.7421
6.9481 | -102.7747
6.9519 | -102.8059
6.9556 | -102.8274
6.9593 | -102.83
6.963 | -102.8148
6.9667 | -102.7919
6.9704 | -102.7676
6.9741 | -102.7366
6.9778 | -102.7033
6.9815 | -102.674
6.9852 | -102.6469
6.9889 | -102.6149
6.9926 | -102.5826
6.9963 | -102.5609
7 | -102.5489
7.0037 | -102.5337
7.0074 | -102.526
7.0111 | -102.527
7.0148 | -102.5399
7.0185 | -102.5653
7.0222 | -102.6038
7.0259 | -102.6574
7.0296 | -102.7269
7.0333 | -102.8061
7.037 | -102.8894
7.0407 | -102.9737
7.0444 | -103.0562
7.0481 | -103.1352
7.0519 | -103.2022
7.0556 | -103.2594
7.0593 | -103.3066
7.063 | -103.3434
7.0667 | -103.3695
7.0704 | -103.3789
7.0741 | -103.3795
7.0778 | -103.3664
7.0815 | -103.3449
7.0852 | -103.3203
7.0889 | -103.284
7.0926 | -103.2455
7.0963 | -103.2135
7.1 | -103.1857
7.1037 | -103.1549
7.1074 | -103.127
7.1111 | -103.1023
7.1148 | -103.0804
7.1185 | -103.0608
7.1222 | -103.0422
7.1259 | -103.0289
7.1296 | -103.0182
7.1333 | -103.0123
7.137 | -103.0096
7.1407 | -103.0152
7.1444 | -103.0269
7.1481 | -103.0437
7.1519 | -103.0654
7.1556 | -103.0873
7.1593 | -103.113
7.163 | -103.1345
7.1667 | -103.1523
7.1704 | -103.1674
7.1741 | -103.1766
7.1778 | -103.183
7.1815 | -103.1888
7.1852 | -103.1869
7.1889 | -103.1783
7.1926 | -103.169
7.1963 | -103.1665
7.2 | -103.1592
7.2037 | -103.1484
7.2074 | -103.1405
7.2111 | -103.1385
7.2148 | -103.1349
7.2185 | -103.1276
7.2222 | -103.1172
7.2259 | -103.1089
7.2296 | -103.1029
7.2333 | -103.0923
7.237 | -103.0771
7.2407 | -103.0629
7.2444 | -103.0529
7.2481 | -103.0412
7.2519 | -103.0281
7.2556 | -103.0214
7.2593 | -103.024
7.263 | -103.0274
7.2667 | -103.0292
7.2704 | -103.034
7.2741 | -103.0376
7.2778 | -103.0432
7.2815 | -103.0513
7.2852 | -103.051
7.2889 | -103.0426
7.2926 | -103.0397
7.2963 | -103.0305
7.3 | -103.0112
7.3037 | -102.9828
7.3074 | -102.9508
7.3111 | -102.921
7.3148 | -102.896
7.3185 | -102.8677
7.3222 | -102.8441
7.3259 | -102.8292
7.3296 | -102.8213
7.3333 | -102.8162
7.337 | -102.8223
7.3407 | -102.8353
7.3444 | -102.8516
7.3481 | -102.8721
7.3519 | -102.892
7.3556 | -102.9081
7.3593 | -102.9264
7.363 | -102.9391
7.3667 | -102.9448
7.3704 | -102.9501
7.3741 | -102.9558
7.3778 | -102.9576
7.3815 | -102.9602
7.3852 | -102.9638
7.3889 | -102.9692
7.3926 | -102.9809
7.3963 | -102.9932
7.4 | -103.0059
7.4037 | -103.0264
7.4074 | -103.0529
7.4111 | -103.0847
7.4148 | -103.1211
7.4185 | -103.1569
7.4222 | -103.1905
7.4259 | -103.2316
7.4296 | -103.2756
7.4333 | -103.3093
7.437 | -103.3383
7.4407 | -103.3684
7.4444 | -103.3881
7.4481 | -103.3999
7.4519 | -103.4083
7.4556 | -103.4042
7.4593 | -103.395
7.463 | -103.3905
7.4667 | -103.3811
7.4704 | -103.3716
7.4741 | -103.3622
7.4778 | -103.3615
7.4815 | -103.3553
7.4852 | -103.3549
7.4889 | -103.3564
7.4926 | -103.3567
7.4963 | -103.3591
7.5 | -103.362
7.5037 | -103.3693
7.5074 | -103.3792
7.5111 | -103.385
7.5148 | -103.3937
7.5185 | -103.4061
7.5222 | -103.4196
7.5259 | -103.4311
7.5296 | -103.4421
7.5333 | -103.4505
7.537 | -103.4551
7.5407 | -103.4612
7.5444 | -103.4728
7.5481 | -103.4788
7.5519 | -103.4789
7.5556 | -103.4778
7.5593 | -103.4751
7.563 | -103.4655
7.5667 | -103.4518
7.5704 | -103.4328
7.5741 | -103.4063
7.5778 | -103.3768
7.5815 | -103.3401
7.5852 | -103.308
7.5889 | -103.2772
7.5926 | -103.2521
7.5963 | -103.2342
7.6 | -103.2222
7.6037 | -103.2179
7.6074 | -103.2177
7.6111 | -103.2229
7.6148 | -103.2308
7.6185 | -103.2365
7.6222 | -103.2427
7.6259 | -103.2494
7.6296 | -103.2485
7.6333 | -103.2473
7.637 | -103.2481
7.6407 | -103.2532
7.6444 | -103.2573
7.6481 | -103.2623
7.6519 | -103.2709
7.6556 | -103.2767
7.6593 | -103.282
7.663 | -103.2961
7.6667 | -103.3143
7.6704 | -103.3346
7.6741 | -103.3612
7.6778 | -103.396
7.6815 | -103.4336
7.6852 | -103.4734
7.6889 | -103.5161
7.6926 | -103.564
7.6963 | -103.6145
7.7 | -103.6566
7.7037 | -103.6952
7.7074 | -103.7294
7.7111 | -103.7544
7.7148 | -103.7685
7.7185 | -103.7797
7.7222 | -103.7878
7.7259 | -103.7831
7.7296 | -103.7769
7.7333 | -103.7668
7.737 | -103.7448
7.7407 | -103.7165
7.7444 | -103.6851
7.7481 | -103.6428
7.7519 | -103.5928
7.7556 | -103.5415
7.7593 | -103.4876
7.763 | -103.4354
7.7667 | -103.3836
7.7704 | -103.3344
7.7741 | -103.2927
7.7778 | -103.2567
7.7815 | -103.234
7.7852 | -103.2227
7.7889 | -103.2235
7.7926 | -103.2346
7.7963 | -103.2587
7.8 | -103.2939
7.8037 | -103.3366
7.8074 | -103.388
7.8111 | -103.4457
7.8148 | -103.5052
7.8185 | -103.5606
7.8222 | -103.6169
7.8259 | -103.6714
7.8296 | -103.7185
7.8333 | -103.7519
7.837 | -103.7805
7.8407 | -103.802
7.8444 | -103.8062
7.8481 | -103.8038
7.8519 | -103.7994
7.8556 | -103.7843
7.8593 | -103.762
7.863 | -103.7428
7.8667 | -103.726
7.8704 | -103.7064
7.8741 | -103.6956
7.8778 | -103.6885
7.8815 | -103.6841
7.8852 | -103.6867
7.8889 | -103.6947
7.8926 | -103.7069
7.8963 | -103.7169
7.9 | -103.7317
7.9037 | -103.7467
7.9074 | -103.757
7.9111 | -103.7641
7.9148 | -103.7702
7.9185 | -103.7767
7.9222 | -103.7778
7.9259 | -103.7814
7.9296 | -103.7851
7.9333 | -103.7899
7.937 | -103.7937
7.9407 | -103.7962
7.9444 | -103.8065
7.9481 | -103.8198
7.9519 | -103.8253
7.9556 | -103.8294
7.9593 | -103.841
7.963 | -103.8494
7.9667 | -103.8539
7.9704 | -103.8621
7.9741 | -103.8717
7.9778 | -103.8752
7.9815 | -103.8769
7.9852 | -103.8825
7.9889 | -103.8817
7.9926 | -103.8791
7.9963 | -103.8754
8 | -103.8713
8.0037 | -103.8641
8.0074 | -103.8534
8.0111 | -103.8421
8.0148 | -103.8359
8.0185 | -103.8277
8.0222 | -103.821
8.0259 | -103.8213
8.0296 | -103.8273
8.0333 | -103.8357
8.037 | -103.8481
8.0407 | -103.8646
8.0444 | -103.8837
8.0481 | -103.9071
8.0519 | -103.9352
8.0556 | -103.9637
8.0593 | -103.9943
8.063 | -104.0208
8.0667 | -104.0493
8.0704 | -104.0788
8.0741 | -104.1036
8.0778 | -104.1226
8.0815 | -104.134
8.0852 | -104.1353
8.0889 | -104.1299
8.0926 | -104.1138
8.0963 | -104.0902
8.1 | -104.0627
8.1037 | -104.0296
8.1074 | -103.9972
8.1111 | -103.9598
8.1148 | -103.9238
8.1185 | -103.8919
8.1222 | -103.8631
8.1259 | -103.835
8.1296 | -103.8117
8.1333 | -103.7918
8.137 | -103.7742
8.1407 | -103.7616
8.1444 | -103.7494
8.1481 | -103.7393
8.1519 | -103.7317
8.1556 | -103.7262
8.1593 | -103.7227
8.163 | -103.721
8.1667 | -103.7191
8.1704 | -103.7202
8.1741 | -103.7178
8.1778 | -103.7154
8.1815 | -103.7153
8.1852 | -103.7187
8.1889 | -103.7218
8.1926 | -103.7259
8.1963 | -103.732
8.2 | -103.7324
8.2037 | -103.7354
8.2074 | -103.7432
8.2111 | -103.7505
8.2148 | -103.7583
8.2185 | -103.7673
8.2222 | -103.7748
8.2259 | -103.7827
8.2296 | -103.7956
8.2333 | -103.8152
8.237 | -103.8328
8.2407 | -103.8455
8.2444 | -103.8626
8.2481 | -103.8864
8.2519 | -103.91
8.2556 | -103.9316
8.2593 | -103.9531
8.263 | -103.973
8.2667 | -103.9977
8.2704 | -104.018
8.2741 | -104.0333
8.2778 | -104.0526
8.2815 | -104.0716
8.2852 | -104.0924
8.2889 | -104.1161
8.2926 | -104.1324
8.2963 | -104.1444
8.3 | -104.1547
8.3037 | -104.1609
8.3074 | -104.1633
8.3111 | -104.1627
8.3148 | -104.1561
8.3185 | -104.1438
8.3222 | -104.1271
8.3259 | -104.1073
8.3296 | -104.0818
8.3333 | -104.0541
8.337 | -104.0321
8.3407 | -104.0121
8.3444 | -103.9952
8.3481 | -103.9789
8.3519 | -103.9674
8.3556 | -103.961
8.3593 | -103.9615
8.363 | -103.9721
8.3667 | -103.984
8.3704 | -103.999
8.3741 | -104.0147
8.3778 | -104.0365
8.3815 | -104.059
8.3852 | -104.0837
8.3889 | -104.1057
8.3926 | -104.1276
8.3963 | -104.1485
8.4 | -104.1636
8.4037 | -104.1754
8.4074 | -104.1837
8.4111 | -104.1852
8.4148 | -104.189
8.4185 | -104.1977
8.4222 | -104.198
8.4259 | -104.1936
8.4296 | -104.1887
8.4333 | -104.1844
8.437 | -104.1834
8.4407 | -104.1831
8.4444 | -104.1864
8.4481 | -104.1921
8.4519 | -104.1981
8.4556 | -104.2062
8.4593 | -104.2101
8.463 | -104.2169
8.4667 | -104.2273
8.4704 | -104.2359
8.4741 | -104.2391
8.4778 | -104.2464
8.4815 | -104.257
8.4852 | -104.2631
8.4889 | -104.2693
8.4926 | -104.273
8.4963 | -104.2805
8.5 | -104.2822
8.5037 | -104.2796
8.5074 | -104.277
8.5111 | -104.2771
8.5148 | -104.2801
8.5185 | -104.2851
8.5222 | -104.2853
8.5259 | -104.2817
8.5296 | -104.2826
8.5333 | -104.2826
8.537 | -104.2819
8.5407 | -104.2776
8.5444 | -104.2757
8.5481 | -104.2692
8.5519 | -104.2647
8.5556 | -104.259
8.5593 | -104.2477
8.563 | -104.2307
8.5667 | -104.2114
8.5704 | -104.1868
8.5741 | -104.1588
8.5778 | -104.1324
8.5815 | -104.1035
8.5852 | -104.0764
8.5889 | -104.0408
8.5926 | -104.0015
8.5963 | -103.9705
8.6 | -103.9442
8.6037 | -103.9163
8.6074 | -103.8852
8.6111 | -103.8582
8.6148 | -103.8385
8.6185 | -103.8273
8.6222 | -103.8171
8.6259 | -103.8105
8.6296 | -103.8128
8.6333 | -103.8174
8.637 | -103.8239
8.6407 | -103.8339
8.6444 | -103.8468
8.6481 | -103.8672
8.6519 | -103.8872
8.6556 | -103.8973
8.6593 | -103.9109
8.663 | -103.9244
8.6667 | -103.9374
8.6704 | -103.9509
8.6741 | -103.9601
8.6778 | -103.9671
8.6815 | -103.9726
8.6852 | -103.9759
8.6889 | -103.9808
8.6926 | -103.9881
8.6963 | -103.9934
8.7 | -103.9975
8.7037 | -104.0025
8.7074 | -104.0064
8.7111 | -104.0114
8.7148 | -104.0226
8.7185 | -104.0479
8.7222 | -104.0759
8.7259 | -104.1014
8.7296 | -104.1265
8.7333 | -104.1581
8.737 | -104.2009
8.7407 | -104.241
8.7444 | -104.2761
8.7481 | -104.3077
8.7519 | -104.3362
8.7556 | -104.3598
8.7593 | -104.3742
8.763 | -104.3788
8.7667 | -104.3757
8.7704 | -104.365
8.7741 | -104.347
8.7778 | -104.3219
8.7815 | -104.2923
8.7852 | -104.2638
8.7889 | -104.2288
8.7926 | -104.1922
8.7963 | -104.1602
8.8 | -104.135
8.8037 | -104.1182
8.8074 | -104.1092
8.8111 | -104.1021
8.8148 | -104.1014
8.8185 | -104.1067
8.8222 | -104.1142
8.8259 | -104.1287
8.8296 | -104.1474
8.8333 | -104.1673
8.837 | -104.1912
8.8407 | -104.2181
8.8444 | -104.2389
8.8481 | -104.2545
8.8519 | -104.2686
8.8556 | -104.2844
8.8593 | -104.2973
8.863 | -104.3082
8.8667 | -104.3124
8.8704 | -104.3172
8.8741 | -104.3182
8.8778 | -104.3171
8.8815 | -104.3118
8.8852 | -104.3019
8.8889 | -104.2879
8.8926 | -104.2772
8.8963 | -104.2677
8.9 | -104.2579
8.9037 | -104.2484
8.9074 | -104.2372
8.9111 | -104.2292
8.9148 | -104.2211
8.9185 | -104.2187
8.9222 | -104.2148
8.9259 | -104.2116
8.9296 | -104.2107
8.9333 | -104.2112
8.937 | -104.2137
8.9407 | -104.2175
8.9444 | -104.2229
8.9481 | -104.2253
8.9519 | -104.2264
8.9556 | -104.223
8.9593 | -104.214
8.963 | -104.2017
8.9667 | -104.1876
8.9704 | -104.1736
8.9741 | -104.1607
8.9778 | -104.1504
8.9815 | -104.1418
8.9852 | -104.1295
8.9889 | -104.1219
8.9926 | -104.1145
8.9963 | -104.108
9 | -104.1054
9.0037 | -104.1056
9.0074 | -104.1078
9.0111 | -104.1128
9.0148 | -104.1222
9.0185 | -104.1277
9.0222 | -104.1281
9.0259 | -104.1262
9.0296 | -104.1275
9.0333 | -104.1323
9.037 | -104.1415
9.0407 | -104.1461
9.0444 | -104.1506
9.0481 | -104.1516
9.0519 | -104.1591
9.0556 | -104.1677
9.0593 | -104.1759
9.063 | -104.1869
9.0667 | -104.2033
9.0704 | -104.2231
9.0741 | -104.2437
9.0778 | -104.265
9.0815 | -104.2877
9.0852 | -104.3111
9.0889 | -104.3318
9.0926 | -104.3534
9.0963 | -104.3737
9.1 | -104.3878
9.1037 | -104.3934
9.1074 | -104.4006
9.1111 | -104.4032
9.1148 | -104.4016
9.1185 | -104.3979
9.1222 | -104.3899
9.1259 | -104.3798
9.1296 | -104.3701
9.1333 | -104.3552
9.137 | -104.3377
9.1407 | -104.3197
9.1444 | -104.3057
9.1481 | -104.2921
9.1519 | -104.2861
9.1556 | -104.284
9.1593 | -104.2816
9.163 | -104.2812
9.1667 | -104.288
9.1704 | -104.2945
9.1741 | -104.3044
9.1778 | -104.3215
9.1815 | -104.3369
9.1852 | -104.3515
9.1889 | -104.3664
9.1926 | -104.382
9.1963 | -104.3962
9.2 | -104.4028
9.2037 | -104.4084
9.2074 | -104.4092
9.2111 | -104.4094
9.2148 | -104.4096
9.2185 | -104.401
9.2222 | -104.3934
9.2259 | -104.3863
9.2296 | -104.3742
9.2333 | -104.3588
9.237 | -104.3454
9.2407 | -104.3348
9.2444 | -104.3213
9.2481 | -104.3071
9.2519 | -104.2964
9.2556 | -104.2897
9.2593 | -104.2833
9.263 | -104.2757
9.2667 | -104.2724
9.2704 | -104.27
9.2741 | -104.2672
9.2778 | -104.2663
9.2815 | -104.266
9.2852 | -104.2661
9.2889 | -104.2704
9.2926 | -104.2779
9.2963 | -104.2905
9.3 | -104.3044
9.3037 | -104.3172
9.3074 | -104.3332
9.3111 | -104.3527
9.3148 | -104.369
9.3185 | -104.39
9.3222 | -104.411
9.3259 | -104.4329
9.3296 | -104.4544
9.3333 | -104.4759
9.337 | -104.4928
9.3407 | -104.5049
9.3444 | -104.5151
9.3481 | -104.52
9.3519 | -104.5188
9.3556 | -104.5135
9.3593 | -104.5009
9.363 | -104.4895
9.3667 | -104.4816
9.3704 | -104.4716
9.3741 | -104.4597
9.3778 | -104.4491
9.3815 | -104.4425
9.3852 | -104.4387
9.3889 | -104.4374
9.3926 | -104.4343
9.3963 | -104.4353
9.4 | -104.4402
9.4037 | -104.4445
9.4074 | -104.4576
9.4111 | -104.4715
9.4148 | -104.4875
9.4185 | -104.5005
9.4222 | -104.5151
9.4259 | -104.5257
9.4296 | -104.5352
9.4333 | -104.5444
9.437 | -104.5535
9.4407 | -104.5576
9.4444 | -104.5598
9.4481 | -104.5564
9.4519 | -104.5524
9.4556 | -104.5476
9.4593 | -104.5358
9.463 | -104.5208
9.4667 | -104.507
9.4704 | -104.4958
9.4741 | -104.4861
9.4778 | -104.4746
9.4815 | -104.4602
9.4852 | -104.4467
9.4889 | -104.4312
9.4926 | -104.415
9.4963 | -104.4002
9.5 | -104.3794
9.5037 | -104.3551
9.5074 | -104.3298
9.5111 | -104.302
9.5148 | -104.2715
9.5185 | -104.2382
9.5222 | -104.2079
9.5259 | -104.182
9.5296 | -104.1606
9.5333 | -104.1404
9.537 | -104.1241
9.5407 | -104.1141
9.5444 | -104.1043
9.5481 | -104.1046
9.5519 | -104.106
9.5556 | -104.1183
9.5593 | -104.1388
9.563 | -104.1644
9.5667 | -104.1918
9.5704 | -104.2219
9.5741 | -104.256
9.5778 | -104.2942
9.5815 | -104.3329
9.5852 | -104.3755
9.5889 | -104.4136
9.5926 | -104.4482
9.5963 | -104.4855
9.6 | -104.5217
9.6037 | -104.5477
9.6074 | -104.5672
9.6111 | -104.5878
9.6148 | -104.5819
9.6185 | -104.5924
9.6222 | -104.6015
9.6259 | -104.6062
9.6296 | -104.6045
9.6333 | -104.5939
9.637 | -104.5836
9.6407 | -104.5757
9.6444 | -104.5702
9.6481 | -104.569
9.6519 | -104.5695
9.6556 | -104.5679
9.6593 | -104.5679
9.663 | -104.5745
9.6667 | -104.5799
9.6704 | -104.5834
9.6741 | -104.5884
9.6778 | -104.5929
9.6815 | -104.5929
9.6852 | -104.5934
9.6889 | -104.5916
9.6926 | -104.5882
9.6963 | -104.5815
9.7 | -104.5735
9.7037 | -104.5642
9.7074 | -104.5571
9.7111 | -104.5474
9.7148 | -104.5317
9.7185 | -104.518
9.7222 | -104.5085
9.7259 | -104.4994
9.7296 | -104.4944
9.7333 | -104.4951
9.737 | -104.4977
9.7407 | -104.5005
9.7444 | -104.511
9.7481 | -104.5285
9.7519 | -104.5477
9.7556 | -104.5742
9.7593 | -104.6042
9.763 | -104.639
9.7667 | -104.6741
9.7704 | -104.7119
9.7741 | -104.7503
9.7778 | -104.789
9.7815 | -104.8209
9.7852 | -104.8516
9.7889 | -104.877
9.7926 | -104.8896
9.7963 | -104.8932
9.8 | -104.9007
9.8037 | -104.9022
9.8074 | -104.8935
9.8111 | -104.8808
9.8148 | -104.8683
9.8185 | -104.8491
9.8222 | -104.8225
9.8259 | -104.7996
9.8296 | -104.7734
9.8333 | -104.7428
9.837 | -104.716
9.8407 | -104.6907
9.8444 | -104.6671
9.8481 | -104.6496
9.8519 | -104.6325
9.8556 | -104.621
9.8593 | -104.6176
9.863 | -104.6157
9.8667 | -104.6163
9.8704 | -104.621
9.8741 | -104.6286
9.8778 | -104.6438
9.8815 | -104.6647
9.8852 | -104.6878
9.8889 | -104.7096
9.8926 | -104.7311
9.8963 | -104.7514
9.9 | -104.771
9.9037 | -104.7863
9.9074 | -104.7943
9.9111 | -104.806
9.9148 | -104.8195
9.9185 | -104.8311
9.9222 | -104.8407
9.9259 | -104.8473
9.9296 | -104.8496
9.9333 | -104.8511
9.937 | -104.8567
9.9407 | -104.8608
9.9444 | -104.8562
9.9481 | -104.8545
9.9519 | -104.8578
9.9556 | -104.8647
9.9593 | -104.8763
9.963 | -104.8933
9.9667 | -104.9133
9.9704 | -104.9321
9.9741 | -104.9606
9.9778 | -104.9964
9.9815 | -105.0324
9.9852 | -105.0679
9.9889 | -105.1039
9.9926 | -105.1393
9.9963 | -105.1783
10 | -105.2129
10.0037 | -105.2405
10.0074 | -105.2651
10.0111 | -105.2862
10.0148 | -105.3006
10.0185 | -105.3104
10.0222 | -105.3088
10.0259 | -105.3025
10.0296 | -105.2982
10.0333 | -105.2888
10.037 | -105.2716
10.0407 | -105.2533
10.0444 | -105.2321
10.0481 | -105.2082
10.0519 | -105.1823
10.0556 | -105.1512
10.0593 | -105.1168
10.063 | -105.0855
10.0667 | -105.0566
10.0704 | -105.0283
10.0741 | -104.999
10.0778 | -104.9713
10.0815 | -104.9426
10.0852 | -104.914
10.0889 | -104.8864
10.0926 | -104.8564
10.0963 | -104.8252
10.1 | -104.7964
10.1037 | -104.7671
10.1074 | -104.7386
10.1111 | -104.7156
10.1148 | -104.6945
10.1185 | -104.6704
10.1222 | -104.6466
10.1259 | -104.626
10.1296 | -104.6115
10.1333 | -104.601
10.137 | -104.5933
10.1407 | -104.5836
10.1444 | -104.5762
10.1481 | -104.5745
10.1519 | -104.5815
10.1556 | -104.5797
10.1593 | -104.5812
10.163 | -104.5889
10.1667 | -104.5947
10.1704 | -104.6043
10.1741 | -104.6162
10.1778 | -104.6238
10.1815 | -104.6344
10.1852 | -104.6443
10.1889 | -104.6532
10.1926 | -104.6633
10.1963 | -104.6743
10.2 | -104.6845
10.2037 | -104.6888
10.2074 | -104.6925
10.2111 | -104.6946
10.2148 | -104.6945
10.2185 | -104.6983
10.2222 | -104.7054
10.2259 | -104.7122
10.2296 | -104.7212
10.2333 | -104.7324
10.237 | -104.7444
10.2407 | -104.7588
10.2444 | -104.7722
10.2481 | -104.7879
10.2519 | -104.8048
10.2556 | -104.8219
10.2593 | -104.8377
10.263 | -104.8519
10.2667 | -104.864
10.2704 | -104.8772
10.2741 | -104.8933
10.2778 | -104.9055
10.2815 | -104.9128
10.2852 | -104.9198
10.2889 | -104.9304
10.2926 | -104.9395
10.2963 | -104.9462
10.3 | -104.9556
10.3037 | -104.9629
10.3074 | -104.9718
10.3111 | -104.9749
10.3148 | -104.9763
10.3185 | -104.9827
10.3222 | -104.9876
10.3259 | -104.9926
10.3296 | -105.0013
10.3333 | -105.0059
10.337 | -105.0124
10.3407 | -105.0246
10.3444 | -105.0339
10.3481 | -105.0407
10.3519 | -105.0499
10.3556 | -105.0625
10.3593 | -105.0725
10.363 | -105.0815
10.3667 | -105.093
10.3704 | -105.1012
10.3741 | -105.1088
10.3778 | -105.1184
10.3815 | -105.1315
10.3852 | -105.144
10.3889 | -105.1516
10.3926 | -105.159
10.3963 | -105.1675
10.4 | -105.1738
10.4037 | -105.1741
10.4074 | -105.171
10.4111 | -105.1658
10.4148 | -105.1591
10.4185 | -105.1514
10.4222 | -105.1393
10.4259 | -105.1191
10.4296 | -105.1018
10.4333 | -105.0851
10.437 | -105.0692
10.4407 | -105.0547
10.4444 | -105.0421
10.4481 | -105.0301
10.4519 | -105.0164
10.4556 | -105.0082
10.4593 | -105.0062
10.463 | -105.0083
10.4667 | -105.0097
10.4704 | -105.0121
10.4741 | -105.0178
10.4778 | -105.0211
10.4815 | -105.0256
10.4852 | -105.0342
10.4889 | -105.0395
10.4926 | -105.0401
10.4963 | -105.0437
10.5 | -105.0478
10.5037 | -105.0484
10.5074 | -105.0489
10.5111 | -105.0511
10.5148 | -105.0511
10.5185 | -105.0496
10.5222 | -105.0416
10.5259 | -105.0335
10.5296 | -105.0228
10.5333 | -105.0095
10.537 | -104.9968
10.5407 | -104.984
10.5444 | -104.9678
10.5481 | -104.9529
10.5519 | -104.9422
10.5556 | -104.9385
10.5593 | -104.9361
10.563 | -104.9329
10.5667 | -104.9297
10.5704 | -104.9252
10.5741 | -104.9203
10.5778 | -104.9208
10.5815 | -104.9177
10.5852 | -104.9134
10.5889 | -104.9079
10.5926 | -104.8987
10.5963 | -104.8869
10.6 | -104.8772
10.6037 | -104.8688
10.6074 | -104.8533
10.6111 | -104.8357
10.6148 | -104.8206
10.6185 | -104.8039
10.6222 | -104.7861
10.6259 | -104.7721
10.6296 | -104.761
10.6333 | -104.7498
10.637 | -104.738
10.6407 | -104.7295
10.6444 | -104.723
10.6481 | -104.7178
10.6519 | -104.7142
10.6556 | -104.7146
10.6593 | -104.7141
10.663 | -104.7167
10.6667 | -104.7224
10.6704 | -104.7249
10.6741 | -104.7326
10.6778 | -104.7426
10.6815 | -104.7546
10.6852 | -104.7673
10.6889 | -104.7765
10.6926 | -104.7847
10.6963 | -104.7989
10.7 | -104.8137
10.7037 | -104.8271
10.7074 | -104.8376
10.7111 | -104.849
10.7148 | -104.8603
10.7185 | -104.8735
10.7222 | -104.8858
10.7259 | -104.8977
10.7296 | -104.907
10.7333 | -104.9101
10.737 | -104.9113
10.7407 | -104.9139
10.7444 | -104.9176
10.7481 | -104.921
10.7519 | -104.9237
10.7556 | -104.9202
10.7593 | -104.9179
10.763 | -104.9177
10.7667 | -104.9159
10.7704 | -104.9124
10.7741 | -104.91
10.7778 | -104.9076
10.7815 | -104.9025
10.7852 | -104.8957
10.7889 | -104.8916
10.7926 | -104.8918
10.7963 | -104.8923
10.8 | -104.8931
10.8037 | -104.8957
10.8074 | -104.9006
10.8111 | -104.9037
10.8148 | -104.9101
10.8185 | -104.9169
10.8222 | -104.9263
10.8259 | -104.9406
10.8296 | -104.9542
10.8333 | -104.9653
10.837 | -104.9742
10.8407 | -104.9826
10.8444 | -104.9907
10.8481 | -104.9986
10.8519 | -105.0059
10.8556 | -105.0072
10.8593 | -105.0064
10.863 | -105.0066
10.8667 | -105.0029
10.8704 | -104.9996
10.8741 | -104.9991
10.8778 | -104.9969
10.8815 | -104.9918
10.8852 | -104.9891
10.8889 | -104.9849
10.8926 | -104.9794
10.8963 | -104.9731
10.9 | -104.9656
10.9037 | -104.9551
10.9074 | -104.9463
10.9111 | -104.9386
10.9148 | -104.9338
10.9185 | -104.9277
10.9222 | -104.922
10.9259 | -104.9203
10.9296 | -104.9196
10.9333 | -104.9192
10.937 | -104.92
10.9407 | -104.9273
10.9444 | -104.9425
10.9481 | -104.9568
10.9519 | -104.9702
10.9556 | -104.9863
10.9593 | -105.0074
10.963 | -105.0326
10.9667 | -105.0588
10.9704 | -105.0891
10.9741 | -105.1177
10.9778 | -105.1457
10.9815 | -105.1689
10.9852 | -105.1884
10.9889 | -105.2096
10.9926 | -105.229
10.9963 | -105.2425
11 | -105.2555
11.0037 | -105.2654
11.0074 | -105.2711
11.0111 | -105.2726
11.0148 | -105.2709
11.0185 | -105.2701
11.0222 | -105.2647
11.0259 | -105.2565
11.0296 | -105.2437
11.0333 | -105.2299
11.037 | -105.2163
11.0407 | -105.2082
11.0444 | -105.1954
11.0481 | -105.1849
11.0519 | -105.1798
11.0556 | -105.1736
11.0593 | -105.1663
11.063 | -105.1647
11.0667 | -105.1639
11.0704 | -105.1659
11.0741 | -105.1721
11.0778 | -105.1807
11.0815 | -105.1913
11.0852 | -105.2019
11.0889 | -105.2182
11.0926 | -105.235
11.0963 | -105.2534
11.1 | -105.2685
11.1037 | -105.2815
11.1074 | -105.2934
11.1111 | -105.3063
11.1148 | -105.3169
11.1185 | -105.3218
11.1222 | -105.3207
11.1259 | -105.3187
11.1296 | -105.314
11.1333 | -105.3074
11.137 | -105.2982
11.1407 | -105.2897
11.1444 | -105.2799
11.1481 | -105.2604
11.1519 | -105.2442
11.1556 | -105.2333
11.1593 | -105.2154
11.163 | -105.1979
11.1667 | -105.1827
11.1704 | -105.1659
11.1741 | -105.156
11.1778 | -105.1475
11.1815 | -105.1443
11.1852 | -105.147
11.1889 | -105.1496
11.1926 | -105.154
11.1963 | -105.1602
11.2 | -105.1631
11.2037 | -105.168
11.2074 | -105.173
11.2111 | -105.1779
11.2148 | -105.1845
11.2185 | -105.18
11.2222 | -105.1808
11.2259 | -105.1752
11.2296 | -105.17
11.2333 | -105.164
11.237 | -105.1536
11.2407 | -105.1393
11.2444 | -105.124
11.2481 | -105.1071
11.2519 | -105.0827
11.2556 | -105.0603
11.2593 | -105.0355
11.263 | -105.0087
11.2667 | -104.9824
11.2704 | -104.9594
11.2741 | -104.9369
11.2778 | -104.9148
11.2815 | -104.8935
11.2852 | -104.8728
11.2889 | -104.8551
11.2926 | -104.8398
11.2963 | -104.829
11.3 | -104.8198
11.3037 | -104.8144
11.3074 | -104.8133
11.3111 | -104.8193
11.3148 | -104.8245
11.3185 | -104.8286
11.3222 | -104.8365
11.3259 | -104.8464
11.3296 | -104.8579
11.3333 | -104.8711
11.337 | -104.8828
11.3407 | -104.8957
11.3444 | -104.908
11.3481 | -104.919
11.3519 | -104.9336
11.3556 | -104.9485
11.3593 | -104.9636
11.363 | -104.9824
11.3667 | -105.0027
11.3704 | -105.0206
11.3741 | -105.0347
11.3778 | -105.0522
11.3815 | -105.0721
11.3852 | -105.0885
11.3889 | -105.1053
11.3926 | -105.1242
11.3963 | -105.1445
11.4 | -105.1644
11.4037 | -105.1836
11.4074 | -105.2019
11.4111 | -105.2181
11.4148 | -105.2337
11.4185 | -105.2501
11.4222 | -105.2589
11.4259 | -105.2676
11.4296 | -105.2749
11.4333 | -105.2785
11.437 | -105.2793
11.4407 | -105.2807
11.4444 | -105.2796
11.4481 | -105.279
11.4519 | -105.2782
11.4556 | -105.2721
11.4593 | -105.2596
11.463 | -105.2456
11.4667 | -105.2348
11.4704 | -105.2249
11.4741 | -105.2127
11.4778 | -105.1978
11.4815 | -105.1821
11.4852 | -105.1688
11.4889 | -105.1566
11.4926 | -105.1438
11.4963 | -105.1256
11.5 | -105.1065
11.5037 | -105.0845
11.5074 | -105.0645
11.5111 | -105.046
11.5148 | -105.0301
11.5185 | -105.0161
11.5222 | -104.9993
11.5259 | -104.9817
11.5296 | -104.9627
11.5333 | -104.9459
11.537 | -104.9323
11.5407 | -104.9204
11.5444 | -104.9089
11.5481 | -104.8978
11.5519 | -104.8914
11.5556 | -104.8876
11.5593 | -104.8811
11.563 | -104.8723
11.5667 | -104.8634
11.5704 | -104.8609
11.5741 | -104.8611
11.5778 | -104.8618
11.5815 | -104.8604
11.5852 | -104.8613
11.5889 | -104.8668
11.5926 | -104.8711
11.5963 | -104.875
11.6 | -104.8808
11.6037 | -104.8902
11.6074 | -104.9011
11.6111 | -104.9132
11.6148 | -104.9266
11.6185 | -104.9449
11.6222 | -104.9655
11.6259 | -104.9884
11.6296 | -105.0086
11.6333 | -105.0298
11.637 | -105.0476
11.6407 | -105.0643
11.6444 | -105.0841
11.6481 | -105.1037
11.6519 | -105.1221
11.6556 | -105.1381
11.6593 | -105.1498
11.663 | -105.1647
11.6667 | -105.1775
11.6704 | -105.1912
11.6741 | -105.2018
11.6778 | -105.2113
11.6815 | -105.2194
11.6852 | -105.2238
11.6889 | -105.2338
11.6926 | -105.247
11.6963 | -105.259
11.7 | -105.2689
11.7037 | -105.2766
11.7074 | -105.2852
11.7111 | -105.2921
11.7148 | -105.2994
11.7185 | -105.3053
11.7222 | -105.3076
11.7259 | -105.3156
11.7296 | -105.3253
11.7333 | -105.332
11.737 | -105.3383
11.7407 | -105.3474
11.7444 | -105.3546
11.7481 | -105.3671
11.7519 | -105.3792
11.7556 | -105.391
11.7593 | -105.4037
11.763 | -105.4166
11.7667 | -105.4339
11.7704 | -105.4502
11.7741 | -105.4658
11.7778 | -105.4773
11.7815 | -105.4937
11.7852 | -105.5098
11.7889 | -105.5211
11.7926 | -105.5352
11.7963 | -105.5499
11.8 | -105.5623
11.8037 | -105.5728
11.8074 | -105.5836
11.8111 | -105.5906
11.8148 | -105.5917
11.8185 | -105.5923
11.8222 | -105.5933
11.8259 | -105.5942
11.8296 | -105.5934
11.8333 | -105.5884
11.837 | -105.5795
11.8407 | -105.5611
11.8444 | -105.5396
11.8481 | -105.5225
11.8519 | -105.5062
11.8556 | -105.4878
11.8593 | -105.4672
11.863 | -105.4441
11.8667 | -105.4214
11.8704 | -105.4007
11.8741 | -105.3805
11.8778 | -105.3612
11.8815 | -105.3466
11.8852 | -105.335
11.8889 | -105.3269
11.8926 | -105.3226
11.8963 | -105.3143
11.9 | -105.3096
11.9037 | -105.3114
11.9074 | -105.3135
11.9111 | -105.3155
11.9148 | -105.324
11.9185 | -105.3341
11.9222 | -105.3461
11.9259 | -105.3584
11.9296 | -105.368
11.9333 | -105.3818
11.937 | -105.3945
11.9407 | -105.4033
11.9444 | -105.415
11.9481 | -105.4244
11.9519 | -105.4301
11.9556 | -105.4296
11.9593 | -105.4297
11.963 | -105.428
11.9667 | -105.4227
11.9704 | -105.4162
11.9741 | -105.4088
11.9778 | -105.3985
11.9815 | -105.3875
11.9852 | -105.3758
11.9889 | -105.364
11.9926 | -105.3483
11.9963 | -105.3293
12 | -105.3121
12.0037 | -105.296
12.0074 | -105.2803
12.0111 | -105.269
12.0148 | -105.262
12.0185 | -105.2553
12.0222 | -105.2418
12.0259 | -105.2285
12.0296 | -105.2195
12.0333 | -105.2119
12.037 | -105.2057
12.0407 | -105.2013
12.0444 | -105.1988
12.0481 | -105.1967
12.0519 | -105.1931
12.0556 | -105.1889
12.0593 | -105.1868
12.063 | -105.1843
12.0667 | -105.1835
12.0704 | -105.1854
12.0741 | -105.1876
12.0778 | -105.1887
12.0815 | -105.1906
12.0852 | -105.192
12.0889 | -105.1911
12.0926 | -105.1901
12.0963 | -105.1916
12.1 | -105.1955
12.1037 | -105.1969
12.1074 | -105.1967
12.1111 | -105.1982
12.1148 | -105.1979
12.1185 | -105.1989
12.1222 | -105.2027
12.1259 | -105.2091
12.1296 | -105.2152
12.1333 | -105.2149
12.137 | -105.22
12.1407 | -105.2266
12.1444 | -105.2328
12.1481 | -105.2391
12.1519 | -105.2425
12.1556 | -105.2435
12.1593 | -105.2445
12.163 | -105.2471
12.1667 | -105.2487
12.1704 | -105.2497
12.1741 | -105.2509
12.1778 | -105.254
12.1815 | -105.2548
12.1852 | -105.2527
12.1889 | -105.2458
12.1926 | -105.2376
12.1963 | -105.2302
12.2 | -105.225
12.2037 | -105.2192
12.2074 | -105.2128
12.2111 | -105.2055
12.2148 | -105.1986
12.2185 | -105.1924
12.2222 | -105.1887
12.2259 | -105.1877
12.2296 | -105.1854
12.2333 | -105.1838
12.237 | -105.1856
12.2407 | -105.1912
12.2444 | -105.1976
12.2481 | -105.2014
12.2519 | -105.2078
12.2556 | -105.2144
12.2593 | -105.2252
12.263 | -105.2384
12.2667 | -105.2507
12.2704 | -105.2636
12.2741 | -105.2738
12.2778 | -105.2853
12.2815 | -105.2928
12.2852 | -105.3005
12.2889 | -105.3071
12.2926 | -105.3195
12.2963 | -105.328
12.3 | -105.3355
12.3037 | -105.3468
12.3074 | -105.3602
12.3111 | -105.3757
12.3148 | -105.387
12.3185 | -105.3988
12.3222 | -105.4093
12.3259 | -105.4224
12.3296 | -105.4376
12.3333 | -105.4565
12.337 | -105.4758
12.3407 | -105.4921
12.3444 | -105.5092
12.3481 | -105.5237
12.3519 | -105.5349
12.3556 | -105.5451
12.3593 | -105.5601
12.363 | -105.5769
12.3667 | -105.5927
12.3704 | -105.6061
12.3741 | -105.619
12.3778 | -105.6294
12.3815 | -105.6383
12.3852 | -105.646
12.3889 | -105.6497
12.3926 | -105.6493
12.3963 | -105.6485
12.4 | -105.6482
12.4037 | -105.6492
12.4074 | -105.6516
12.4111 | -105.6505
12.4148 | -105.6468
12.4185 | -105.6442
12.4222 | -105.6418
12.4259 | -105.6384
12.4296 | -105.6338
12.4333 | -105.6292
12.437 | -105.62
12.4407 | -105.6051
12.4444 | -105.5915
12.4481 | -105.5794
12.4519 | -105.5655
12.4556 | -105.5516
12.4593 | -105.5346
12.463 | -105.5139
12.4667 | -105.4945
12.4704 | -105.4736
12.4741 | -105.4483
12.4778 | -105.4223
12.4815 | -105.3953
12.4852 | -105.3662
12.4889 | -105.3382
12.4926 | -105.3106
12.4963 | -105.2811
12.5 | -105.2527
12.5037 | -105.2246
12.5074 | -105.197
12.5111 | -105.172
12.5148 | -105.1461
12.5185 | -105.1215
12.5222 | -105.1004
12.5259 | -105.0815
12.5296 | -105.0682
12.5333 | -105.0578
12.537 | -105.0498
12.5407 | -105.0454
12.5444 | -105.0422
12.5481 | -105.0466
12.5519 | -105.0553
12.5556 | -105.0667
12.5593 | -105.08
12.563 | -105.0956
12.5667 | -105.1122
12.5704 | -105.1344
12.5741 | -105.1549
12.5778 | -105.1768
12.5815 | -105.198
12.5852 | -105.2152
12.5889 | -105.2342
12.5926 | -105.2567
12.5963 | -105.2797
12.6 | -105.2965
12.6037 | -105.3143
12.6074 | -105.3344
12.6111 | -105.3506
12.6148 | -105.3661
12.6185 | -105.3829
12.6222 | -105.4006
12.6259 | -105.416
12.6296 | -105.4274
12.6333 | -105.4427
12.637 | -105.4558
12.6407 | -105.4686
12.6444 | -105.477
12.6481 | -105.4818
12.6519 | -105.4901
12.6556 | -105.5001
12.6593 | -105.5108
12.663 | -105.5207
12.6667 | -105.5302
12.6704 | -105.5366
12.6741 | -105.5427
12.6778 | -105.5505
12.6815 | -105.5575
12.6852 | -105.5679
12.6889 | -105.5765
12.6926 | -105.5809
12.6963 | -105.5864
12.7 | -105.5873
12.7037 | -105.5853
12.7074 | -105.5826
12.7111 | -105.5797
12.7148 | -105.5773
12.7185 | -105.5731
12.7222 | -105.5634
12.7259 | -105.5552
12.7296 | -105.5462
12.7333 | -105.5334
12.737 | -105.5187
12.7407 | -105.5027
12.7444 | -105.489
12.7481 | -105.4781
12.7519 | -105.4655
12.7556 | -105.4543
12.7593 | -105.4394
12.763 | -105.428
12.7667 | -105.4198
12.7704 | -105.407
12.7741 | -105.3957
12.7778 | -105.3886
12.7815 | -105.3855
12.7852 | -105.3859
12.7889 | -105.3857
12.7926 | -105.3882
12.7963 | -105.3935
12.8 | -105.4002
12.8037 | -105.4079
12.8074 | -105.4157
12.8111 | -105.4253
12.8148 | -105.4339
12.8185 | -105.4445
12.8222 | -105.4566
12.8259 | -105.4669
12.8296 | -105.4781
12.8333 | -105.4881
12.837 | -105.4937
12.8407 | -105.4991
12.8444 | -105.5057
12.8481 | -105.5107
12.8519 | -105.5155
12.8556 | -105.5217
12.8593 | -105.5254
12.863 | -105.5287
12.8667 | -105.5326
12.8704 | -105.5341
12.8741 | -105.5333
12.8778 | -105.5334
12.8815 | -105.5326
12.8852 | -105.5334
12.8889 | -105.5323
12.8926 | -105.5274
12.8963 | -105.5177
12.9 | -105.5041
12.9037 | -105.4914
12.9074 | -105.4822
12.9111 | -105.4706
12.9148 | -105.4557
12.9185 | -105.4441
12.9222 | -105.4354
12.9259 | -105.4261
12.9296 | -105.4123
12.9333 | -105.3993
12.937 | -105.388
12.9407 | -105.3777
12.9444 | -105.371
12.9481 | -105.3641
12.9519 | -105.3588
12.9556 | -105.3581
12.9593 | -105.3575
12.963 | -105.3578
12.9667 | -105.3642
12.9704 | -105.3701
12.9741 | -105.3823
12.9778 | -105.3957
12.9815 | -105.4042
12.9852 | -105.4134
12.9889 | -105.4268
12.9926 | -105.4414
12.9963 | -105.4567
13 | -105.473
13.0037 | -105.4891
13.0074 | -105.5051
13.0111 | -105.5226
13.0148 | -105.5415
13.0185 | -105.5628
13.0222 | -105.5768
13.0259 | -105.5894
13.0296 | -105.5984
13.0333 | -105.6074
13.037 | -105.6165
13.0407 | -105.6202
13.0444 | -105.6233
13.0481 | -105.6272
13.0519 | -105.6247
13.0556 | -105.6132
13.0593 | -105.6007
13.063 | -105.5909
13.0667 | -105.5799
13.0704 | -105.5685
13.0741 | -105.557
13.0778 | -105.5414
13.0815 | -105.5258
13.0852 | -105.5105
13.0889 | -105.4943
13.0926 | -105.4781
13.0963 | -105.4599
13.1 | -105.4438
13.1037 | -105.4283
13.1074 | -105.412
13.1111 | -105.3947
13.1148 | -105.3807
13.1185 | -105.3696
13.1222 | -105.36
13.1259 | -105.3562
13.1296 | -105.3562
13.1333 | -105.3551
13.137 | -105.357
13.1407 | -105.363
13.1444 | -105.3698
13.1481 | -105.3792
13.1519 | -105.3889
13.1556 | -105.397
13.1593 | -105.4113
13.163 | -105.4213
13.1667 | -105.4342
13.1704 | -105.4532
13.1741 | -105.4723
13.1778 | -105.4906
13.1815 | -105.509
13.1852 | -105.5249
13.1889 | -105.5374
13.1926 | -105.5466
13.1963 | -105.5525
13.2 | -105.5553
13.2037 | -105.5548
13.2074 | -105.5521
13.2111 | -105.5539
13.2148 | -105.5513
13.2185 | -105.5482
13.2222 | -105.5395
13.2259 | -105.5303
13.2296 | -105.5229
13.2333 | -105.5136
13.237 | -105.5006
13.2407 | -105.4908
13.2444 | -105.4837
13.2481 | -105.4748
13.2519 | -105.464
13.2556 | -105.4532
13.2593 | -105.4424
13.263 | -105.4319
13.2667 | -105.4254
13.2704 | -105.4232
13.2741 | -105.4207
13.2778 | -105.4204
13.2815 | -105.421
13.2852 | -105.4188
13.2889 | -105.4165
13.2926 | -105.4201
13.2963 | -105.4256
13.3 | -105.4322
13.3037 | -105.4369
13.3074 | -105.4411
13.3111 | -105.4485
13.3148 | -105.4529
13.3185 | -105.4547
13.3222 | -105.4579
13.3259 | -105.4579
13.3296 | -105.4571
13.3333 | -105.4609
13.337 | -105.4624
13.3407 | -105.4626
13.3444 | -105.4642
13.3481 | -105.4602
13.3519 | -105.4572
13.3556 | -105.4364
13.3593 | -105.4231
13.363 | -105.4104
13.3667 | -105.3988
13.3704 | -105.383
13.3741 | -105.3688
13.3778 | -105.3515
13.3815 | -105.3317
13.3852 | -105.3136
13.3889 | -105.2953
13.3926 | -105.2783
13.3963 | -105.262
13.4 | -105.2436
13.4037 | -105.2234
13.4074 | -105.2052
13.4111 | -105.1865
13.4148 | -105.166
13.4185 | -105.1468
13.4222 | -105.1302
13.4259 | -105.1133
13.4296 | -105.1005
13.4333 | -105.0897
13.437 | -105.0808
13.4407 | -105.0734
13.4444 | -105.065
13.4481 | -105.059
13.4519 | -105.0552
13.4556 | -105.0544
13.4593 | -105.0562
13.463 | -105.0565
13.4667 | -105.0598
13.4704 | -105.0637
13.4741 | -105.0691
13.4778 | -105.0809
13.4815 | -105.0926
13.4852 | -105.1031
13.4889 | -105.1122
13.4926 | -105.1247
13.4963 | -105.1416
13.5 | -105.1602
13.5037 | -105.1751
13.5074 | -105.1867
13.5111 | -105.2007
13.5148 | -105.2159
13.5185 | -105.2299
13.5222 | -105.244
13.5259 | -105.2521
13.5296 | -105.2604
13.5333 | -105.2645
13.537 | -105.2653
13.5407 | -105.2701
13.5444 | -105.2765
13.5481 | -105.2789
13.5519 | -105.281
13.5556 | -105.2829
13.5593 | -105.2793
13.563 | -105.271
13.5667 | -105.2642
13.5704 | -105.2561
13.5741 | -105.2517
13.5778 | -105.2504
13.5815 | -105.2445
13.5852 | -105.238
13.5889 | -105.2321
13.5926 | -105.2242
13.5963 | -105.2166
13.6 | -105.2125
13.6037 | -105.2133
13.6074 | -105.2135
13.6111 | -105.214
13.6148 | -105.2166
13.6185 | -105.2169
13.6222 | -105.2172
13.6259 | -105.2186
13.6296 | -105.2219
13.6333 | -105.2274
13.637 | -105.2346
13.6407 | -105.2424
13.6444 | -105.2466
13.6481 | -105.2523
13.6519 | -105.2591
13.6556 | -105.2647
13.6593 | -105.2724
13.663 | -105.2809
13.6667 | -105.2872
13.6704 | -105.2902
13.6741 | -105.2922
13.6778 | -105.2949
13.6815 | -105.295
13.6852 | -105.297
13.6889 | -105.2961
13.6926 | -105.2911
13.6963 | -105.2874
13.7 | -105.2833
13.7037 | -105.2807
13.7074 | -105.278
13.7111 | -105.2699
13.7148 | -105.2681
13.7185 | -105.2654
13.7222 | -105.2628
13.7259 | -105.2613
13.7296 | -105.2636
13.7333 | -105.2658
13.737 | -105.2659
13.7407 | -105.2635
13.7444 | -105.267
13.7481 | -105.2702
13.7519 | -105.2739
13.7556 | -105.2766
13.7593 | -105.2822
13.763 | -105.2915
13.7667 | -105.3007
13.7704 | -105.3096
13.7741 | -105.3175
13.7778 | -105.33
13.7815 | -105.3443
13.7852 | -105.356
13.7889 | -105.3692
13.7926 | -105.388
13.7963 | -105.4062
13.8 | -105.4228
13.8037 | -105.4389
13.8074 | -105.4599
13.8111 | -105.4779
13.8148 | -105.4975
13.8185 | -105.521
13.8222 | -105.5412
13.8259 | -105.5591
13.8296 | -105.5771
13.8333 | -105.5968
13.837 | -105.6173
13.8407 | -105.6386
13.8444 | -105.6596
13.8481 | -105.68
13.8519 | -105.7014
13.8556 | -105.7218
13.8593 | -105.7406
13.863 | -105.7639
13.8667 | -105.7849
13.8704 | -105.802
13.8741 | -105.819
13.8778 | -105.8374
13.8815 | -105.8551
13.8852 | -105.8687
13.8889 | -105.8764
13.8926 | -105.889
13.8963 | -105.8972
13.9 | -105.9065
13.9037 | -105.9176
13.9074 | -105.9255
13.9111 | -105.9312
13.9148 | -105.9306
13.9185 | -105.9284
13.9222 | -105.9308
13.9259 | -105.9315
13.9296 | -105.9314
13.9333 | -105.9318
13.937 | -105.9307
13.9407 | -105.9284
13.9444 | -105.9241
13.9481 | -105.9185
13.9519 | -105.9121
13.9556 | -105.905
13.9593 | -105.8901
13.963 | -105.8646
13.9667 | -105.8426
13.9704 | -105.8258
13.9741 | -105.8081
13.9778 | -105.7836
13.9815 | -105.7563
13.9852 | -105.7301
13.9889 | -105.7027
13.9926 | -105.6731
13.9963 | -105.643
14 | -105.6089
14.0037 | -105.5753
14.0074 | -105.5403
14.0111 | -105.5062
14.0148 | -105.4743
14.0185 | -105.44
14.0222 | -105.4064
14.0259 | -105.3738
14.0296 | -105.341
14.0333 | -105.3096
14.037 | -105.2824
14.0407 | -105.2591
14.0444 | -105.2352
14.0481 | -105.2119
14.0519 | -105.1955
14.0556 | -105.1824
14.0593 | -105.1733
14.063 | -105.1688
14.0667 | -105.171
14.0704 | -105.1768
14.0741 | -105.1825
14.0778 | -105.1932
14.0815 | -105.2106
14.0852 | -105.2329
14.0889 | -105.2574
14.0926 | -105.2807
14.0963 | -105.3073
14.1 | -105.3343
14.1037 | -105.3619
14.1074 | -105.3915
14.1111 | -105.4165
14.1148 | -105.4433
14.1185 | -105.4714
14.1222 | -105.4981
14.1259 | -105.5175
14.1296 | -105.5343
14.1333 | -105.5503
14.137 | -105.5656
14.1407 | -105.5799
14.1444 | -105.5942
14.1481 | -105.6069
14.1519 | -105.6176
14.1556 | -105.625
14.1593 | -105.6316
14.163 | -105.642
14.1667 | -105.6525
14.1704 | -105.6569
14.1741 | -105.6557
14.1778 | -105.6569
14.1815 | -105.6591
14.1852 | -105.6634
14.1889 | -105.6679
14.1926 | -105.6696
14.1963 | -105.6694
14.2 | -105.6718
14.2037 | -105.6791
14.2074 | -105.6872
14.2111 | -105.6925
14.2148 | -105.6966
14.2185 | -105.7015
14.2222 | -105.7079
14.2259 | -105.7161
14.2296 | -105.723
14.2333 | -105.7334
14.237 | -105.7417
14.2407 | -105.7489
14.2444 | -105.7557
14.2481 | -105.7613
14.2519 | -105.7661
14.2556 | -105.7648
14.2593 | -105.7662
14.263 | -105.7673
14.2667 | -105.7662
14.2704 | -105.7655
14.2741 | -105.7651
14.2778 | -105.7675
14.2815 | -105.7657
14.2852 | -105.7619
14.2889 | -105.7586
14.2926 | -105.7573
14.2963 | -105.7577
14.3 | -105.7561
14.3037 | -105.7562
14.3074 | -105.7549
14.3111 | -105.7504
14.3148 | -105.7537
14.3185 | -105.7566
14.3222 | -105.7567
14.3259 | -105.7572
14.3296 | -105.7605
14.3333 | -105.7648
14.337 | -105.77
14.3407 | -105.7748
14.3444 | -105.7793
14.3481 | -105.7841
14.3519 | -105.7913
14.3556 | -105.8018
14.3593 | -105.8129
14.363 | -105.826
14.3667 | -105.8395
14.3704 | -105.8545
14.3741 | -105.8676
14.3778 | -105.879
14.3815 | -105.8896
14.3852 | -105.8989
14.3889 | -105.9061
14.3926 | -105.9153
14.3963 | -105.9277
14.4 | -105.9384
14.4037 | -105.9425
14.4074 | -105.9486
14.4111 | -105.9556
14.4148 | -105.9602
14.4185 | -105.9604
14.4222 | -105.9521
14.4259 | -105.9266
14.4296 | -105.9178
14.4333 | -105.905
14.437 | -105.8943
14.4407 | -105.8786
14.4444 | -105.8618
14.4481 | -105.844
14.4519 | -105.8267
14.4556 | -105.8085
14.4593 | -105.7876
14.463 | -105.7653
14.4667 | -105.7406
14.4704 | -105.7145
14.4741 | -105.6855
14.4778 | -105.6596
14.4815 | -105.6357
14.4852 | -105.612
14.4889 | -105.5906
14.4926 | -105.5681
14.4963 | -105.548
14.5 | -105.5277
14.5037 | -105.5123
14.5074 | -105.5001
14.5111 | -105.4881
14.5148 | -105.4778
14.5185 | -105.4711
14.5222 | -105.4635
14.5259 | -105.4591
14.5296 | -105.4585
14.5333 | -105.4585
14.537 | -105.4595
14.5407 | -105.464
14.5444 | -105.4703
14.5481 | -105.4791
14.5519 | -105.4917
14.5556 | -105.5056
14.5593 | -105.5202
14.563 | -105.5341
14.5667 | -105.5492
14.5704 | -105.5658
14.5741 | -105.5831
14.5778 | -105.6021
14.5815 | -105.6224
14.5852 | -105.645
14.5889 | -105.6668
14.5926 | -105.6874
14.5963 | -105.7065
14.6 | -105.7241
14.6037 | -105.7393
14.6074 | -105.7557
14.6111 | -105.7745
14.6148 | -105.7885
14.6185 | -105.8016
14.6222 | -105.8167
14.6259 | -105.8305
14.6296 | -105.8476
14.6333 | -105.8645
14.637 | -105.8741
14.6407 | -105.8828
14.6444 | -105.8925
14.6481 | -105.8998
14.6519 | -105.9071
14.6556 | -105.9179
14.6593 | -105.9276
14.663 | -105.9344
14.6667 | -105.9406
14.6704 | -105.9498
14.6741 | -105.9604
14.6778 | -105.9654
14.6815 | -105.9724
14.6852 | -105.9803
14.6889 | -105.9793
14.6926 | -105.9772
14.6963 | -105.975
14.7 | -105.9702
14.7037 | -105.9648
14.7074 | -105.9581
14.7111 | -105.9514
14.7148 | -105.9396
14.7185 | -105.9271
14.7222 | -105.9133
14.7259 | -105.8958
14.7296 | -105.8782
14.7333 | -105.8585
14.737 | -105.8367
14.7407 | -105.8147
14.7444 | -105.792
14.7481 | -105.7713
14.7519 | -105.748
14.7556 | -105.7255
14.7593 | -105.7011
14.763 | -105.6772
14.7667 | -105.656
14.7704 | -105.6351
14.7741 | -105.6165
14.7778 | -105.5987
14.7815 | -105.5819
14.7852 | -105.5673
14.7889 | -105.5524
14.7926 | -105.5389
14.7963 | -105.5274
14.8 | -105.518
14.8037 | -105.5126
14.8074 | -105.5094
14.8111 | -105.506
14.8148 | -105.5057
14.8185 | -105.5064
14.8222 | -105.5095
14.8259 | -105.516
14.8296 | -105.5221
14.8333 | -105.5301
14.837 | -105.5428
14.8407 | -105.5581
14.8444 | -105.5751
14.8481 | -105.5909
14.8519 | -105.6089
14.8556 | -105.629
14.8593 | -105.6482
14.863 | -105.6656
14.8667 | -105.6813
14.8704 | -105.6971
14.8741 | -105.7176
14.8778 | -105.7386
14.8815 | -105.756
14.8852 | -105.7695
14.8889 | -105.786
14.8926 | -105.7997
14.8963 | -105.8113
14.9 | -105.822
14.9037 | -105.8294
14.9074 | -105.8313
14.9111 | -105.8343
14.9148 | -105.8365
14.9185 | -105.8361
14.9222 | -105.8354
14.9259 | -105.8333
14.9296 | -105.8293
14.9333 | -105.8265
14.937 | -105.8238
14.9407 | -105.8172
14.9444 | -105.8129
14.9481 | -105.8087
14.9519 | -105.8043
14.9556 | -105.7992
14.9593 | -105.7938
14.963 | -105.7889
14.9667 | -105.7859
14.9704 | -105.7823
14.9741 | -105.7807
14.9778 | -105.778
14.9815 | -105.7725
14.9852 | -105.7682
14.9889 | -105.7647
14.9926 | -105.7618
14.9963 | -105.7592
15 | -105.7581
15.0037 | -105.7568
15.0074 | -105.7563
15.0111 | -105.7549
15.0148 | -105.7512
15.0185 | -105.7467
15.0222 | -105.7435
15.0259 | -105.739
15.0296 | -105.7308
15.0333 | -105.7219
15.037 | -105.7135
15.0407 | -105.7051
15.0444 | -105.6992
15.0481 | -105.6893
15.0519 | -105.6757
15.0556 | -105.6613
15.0593 | -105.644
15.063 | -105.6279
15.0667 | -105.611
15.0704 | -105.5922
15.0741 | -105.574
15.0778 | -105.5574
15.0815 | -105.5405
15.0852 | -105.5215
15.0889 | -105.4994
15.0926 | -105.4784
15.0963 | -105.4581
15.1 | -105.441
15.1037 | -105.4246
15.1074 | -105.4084
15.1111 | -105.3946
15.1148 | -105.3839
15.1185 | -105.3726
15.1222 | -105.3669
15.1259 | -105.3612
15.1296 | -105.3558
15.1333 | -105.3541
15.137 | -105.3533
15.1407 | -105.3573
15.1444 | -105.3615
15.1481 | -105.3662
15.1519 | -105.3734
15.1556 | -105.3848
15.1593 | -105.3986
15.163 | -105.4121
15.1667 | -105.4254
15.1704 | -105.444
15.1741 | -105.4613
15.1778 | -105.4808
15.1815 | -105.5011
15.1852 | -105.5197
15.1889 | -105.5399
15.1926 | -105.5607
15.1963 | -105.5794
15.2 | -105.6006
15.2037 | -105.6208
15.2074 | -105.6403
15.2111 | -105.6581
15.2148 | -105.6737
15.2185 | -105.6873
15.2222 | -105.7002
15.2259 | -105.7145
15.2296 | -105.7238
15.2333 | -105.732
15.237 | -105.7413
15.2407 | -105.7469
15.2444 | -105.7532
15.2481 | -105.7599
15.2519 | -105.7645
15.2556 | -105.7642
15.2593 | -105.7625
15.263 | -105.7596
15.2667 | -105.7579
15.2704 | -105.7551
15.2741 | -105.7507
15.2778 | -105.7493
15.2815 | -105.7468
15.2852 | -105.742
15.2889 | -105.7372
15.2926 | -105.7307
15.2963 | -105.7251
15.3 | -105.7198
15.3037 | -105.712
15.3074 | -105.7016
15.3111 | -105.6914
15.3148 | -105.6843
15.3185 | -105.6756
15.3222 | -105.6646
15.3259 | -105.6551
15.3296 | -105.6459
15.3333 | -105.6347
15.337 | -105.6235
15.3407 | -105.6092
15.3444 | -105.5931
15.3481 | -105.5805
15.3519 | -105.5657
15.3556 | -105.549
15.3593 | -105.5315
15.363 | -105.5158
15.3667 | -105.4974
15.3704 | -105.4775
15.3741 | -105.457
15.3778 | -105.4341
15.3815 | -105.4111
15.3852 | -105.3896
15.3889 | -105.3672
15.3926 | -105.3472
15.3963 | -105.3259
15.4 | -105.3036
15.4037 | -105.2833
15.4074 | -105.2634
15.4111 | -105.2433
15.4148 | -105.2263
15.4185 | -105.2104
15.4222 | -105.1943
15.4259 | -105.1796
15.4296 | -105.1685
15.4333 | -105.1601
15.437 | -105.152
15.4407 | -105.1478
15.4444 | -105.1415
15.4481 | -105.1395
15.4519 | -105.1396
15.4556 | -105.1386
15.4593 | -105.1378
15.463 | -105.1396
15.4667 | -105.142
15.4704 | -105.1455
15.4741 | -105.1529
15.4778 | -105.1585
15.4815 | -105.1656
15.4852 | -105.1735
15.4889 | -105.184
15.4926 | -105.1998
15.4963 | -105.2103
15.5 | -105.221
15.5037 | -105.2302
15.5074 | -105.2366
15.5111 | -105.2459
15.5148 | -105.2577
15.5185 | -105.2681
15.5222 | -105.2785
15.5259 | -105.2916
15.5296 | -105.3041
15.5333 | -105.3146
15.537 | -105.3232
15.5407 | -105.3276
15.5444 | -105.332
15.5481 | -105.3346
15.5519 | -105.3363
15.5556 | -105.3393
15.5593 | -105.3407
15.563 | -105.3408
15.5667 | -105.3384
15.5704 | -105.3332
15.5741 | -105.3267
15.5778 | -105.3176
15.5815 | -105.3071
15.5852 | -105.2984
15.5889 | -105.2907
15.5926 | -105.2799
15.5963 | -105.2658
15.6 | -105.2497
15.6037 | -105.2361
15.6074 | -105.2193
15.6111 | -105.2004
15.6148 | -105.1824
15.6185 | -105.1619
15.6222 | -105.1422
15.6259 | -105.1243
15.6296 | -105.1055
15.6333 | -105.0852
15.637 | -105.0649
15.6407 | -105.0446
15.6444 | -105.0238
15.6481 | -105.003
15.6519 | -104.9794
15.6556 | -104.9552
15.6593 | -104.9346
15.663 | -104.9124
15.6667 | -104.8919
15.6704 | -104.8732
15.6741 | -104.8516
15.6778 | -104.8313
15.6815 | -104.8135
15.6852 | -104.7942
15.6889 | -104.7749
15.6926 | -104.7545
15.6963 | -104.7367
15.7 | -104.7187
15.7037 | -104.6985
15.7074 | -104.6809
15.7111 | -104.6645
15.7148 | -104.6471
15.7185 | -104.6308
15.7222 | -104.6175
15.7259 | -104.6048
15.7296 | -104.5932
15.7333 | -104.5831
15.737 | -104.572
15.7407 | -104.5612
15.7444 | -104.552
15.7481 | -104.5454
15.7519 | -104.5408
15.7556 | -104.5377
15.7593 | -104.5358
15.763 | -104.5362
15.7667 | -104.5401
15.7704 | -104.5472
15.7741 | -104.557
15.7778 | -104.5674
15.7815 | -104.5796
15.7852 | -104.5957
15.7889 | -104.6113
15.7926 | -104.6284
15.7963 | -104.6454
15.8 | -104.6635
15.8037 | -104.6834
15.8074 | -104.7014
15.8111 | -104.7181
15.8148 | -104.7365
15.8185 | -104.7564
15.8222 | -104.7764
15.8259 | -104.7985
15.8296 | -104.8196
15.8333 | -104.84
15.837 | -104.8579
15.8407 | -104.8745
15.8444 | -104.892
15.8481 | -104.908
15.8519 | -104.9209
15.8556 | -104.9319
15.8593 | -104.945
15.863 | -104.9589
15.8667 | -104.9722
15.8704 | -104.9845
15.8741 | -104.9947
15.8778 | -105.0053
15.8815 | -105.0161
15.8852 | -105.0237
15.8889 | -105.0322
15.8926 | -105.0442
15.8963 | -105.0559
15.9 | -105.0693
15.9037 | -105.0795
15.9074 | -105.0887
15.9111 | -105.0995
15.9148 | -105.1093
15.9185 | -105.1244
15.9222 | -105.1363
15.9259 | -105.1515
15.9296 | -105.1682
15.9333 | -105.1841
15.937 | -105.2002
15.9407 | -105.2182
15.9444 | -105.235
15.9481 | -105.2512
15.9519 | -105.2661
15.9556 | -105.2819
15.9593 | -105.3014
15.963 | -105.3187
15.9667 | -105.334
15.9704 | -105.3464
15.9741 | -105.3568
15.9778 | -105.3664
15.9815 | -105.3776
15.9852 | -105.3905
15.9889 | -105.4005
15.9926 | -105.4091
15.9963 | -105.4168
16 | -105.418
16.0037 | -105.4202
16.0074 | -105.4228
16.0111 | -105.4256
16.0148 | -105.4313
16.0185 | -105.4355
16.0222 | -105.4379
16.0259 | -105.44
16.0296 | -105.443
16.0333 | -105.4466
16.037 | -105.4496
16.0407 | -105.4504
16.0444 | -105.4508
16.0481 | -105.4533
16.0519 | -105.4558
16.0556 | -105.4553
16.0593 | -105.4559
16.063 | -105.4615
16.0667 | -105.4696
16.0704 | -105.4773
16.0741 | -105.4844
16.0778 | -105.4896
16.0815 | -105.4971
16.0852 | -105.5078
16.0889 | -105.52
16.0926 | -105.5298
16.0963 | -105.5383
16.1 | -105.549
16.1037 | -105.5612
16.1074 | -105.5752
16.1111 | -105.5873
16.1148 | -105.5998
16.1185 | -105.613
16.1222 | -105.6212
16.1259 | -105.6302
16.1296 | -105.6395
16.1333 | -105.6498
16.137 | -105.6588
16.1407 | -105.6666
16.1444 | -105.6762
16.1481 | -105.6802
16.1519 | -105.6819
16.1556 | -105.6815
16.1593 | -105.6814
16.163 | -105.683
16.1667 | -105.6839
16.1704 | -105.6823
16.1741 | -105.6841
16.1778 | -105.6869
16.1815 | -105.6883
16.1852 | -105.6894
16.1889 | -105.6882
16.1926 | -105.688
16.1963 | -105.69
16.2 | -105.6941
16.2037 | -105.6961
16.2074 | -105.6965
16.2111 | -105.6979
16.2148 | -105.6945
16.2185 | -105.6922
16.2222 | -105.6933
16.2259 | -105.6929
16.2296 | -105.6952
16.2333 | -105.6952
16.237 | -105.6948
16.2407 | -105.6917
16.2444 | -105.6865
16.2481 | -105.6837
16.2519 | -105.6804
16.2556 | -105.677
16.2593 | -105.673
16.263 | -105.668
16.2667 | -105.6613
16.2704 | -105.6542
16.2741 | -105.6489
16.2778 | -105.6404
16.2815 | -105.63
16.2852 | -105.6197
16.2889 | -105.6126
16.2926 | -105.601
16.2963 | -105.5867
16.3 | -105.5735
16.3037 | -105.56
16.3074 | -105.5451
16.3111 | -105.5315
16.3148 | -105.5166
16.3185 | -105.4996
16.3222 | -105.4802
16.3259 | -105.4632
16.3296 | -105.4487
16.3333 | -105.4342
16.337 | -105.4194
16.3407 | -105.4049
16.3444 | -105.3916
16.3481 | -105.3794
16.3519 | -105.3687
16.3556 | -105.3599
16.3593 | -105.3526
16.363 | -105.3472
16.3667 | -105.3437
16.3704 | -105.3405
16.3741 | -105.3376
16.3778 | -105.3368
16.3815 | -105.3355
16.3852 | -105.3391
16.3889 | -105.3457
16.3926 | -105.3538
16.3963 | -105.3606
16.4 | -105.3675
16.4037 | -105.3768
16.4074 | -105.3843
16.4111 | -105.395
16.4148 | -105.4079
16.4185 | -105.4209
16.4222 | -105.4354
16.4259 | -105.4476
16.4296 | -105.4583
16.4333 | -105.4711
16.437 | -105.486
16.4407 | -105.5025
16.4444 | -105.517
16.4481 | -105.5327
16.4519 | -105.5478
16.4556 | -105.5649
16.4593 | -105.5831
16.463 | -105.6013
16.4667 | -105.6189
16.4704 | -105.6367
16.4741 | -105.6565
16.4778 | -105.6766
16.4815 | -105.6943
16.4852 | -105.7128
16.4889 | -105.7316
16.4926 | -105.7494
16.4963 | -105.765
16.5 | -105.7782
16.5037 | -105.7925
16.5074 | -105.804
16.5111 | -105.816
16.5148 | -105.8254
16.5185 | -105.833
16.5222 | -105.8374
16.5259 | -105.8425
16.5296 | -105.8464
16.5333 | -105.848
16.537 | -105.8469
16.5407 | -105.8474
16.5444 | -105.845
16.5481 | -105.8405
16.5519 | -105.8343
16.5556 | -105.8272
16.5593 | -105.8171
16.563 | -105.8022
16.5667 | -105.7925
16.5704 | -105.7811
16.5741 | -105.7703
16.5778 | -105.7585
16.5815 | -105.7486
16.5852 | -105.7365
16.5889 | -105.7238
16.5926 | -105.7124
16.5963 | -105.7019
16.6 | -105.6947
16.6037 | -105.6886
16.6074 | -105.6838
16.6111 | -105.6816
16.6148 | -105.6791
16.6185 | -105.6797
16.6222 | -105.6754
16.6259 | -105.6724
16.6296 | -105.6703
16.6333 | -105.6697
16.637 | -105.6726
16.6407 | -105.678
16.6444 | -105.6852
16.6481 | -105.6893
16.6519 | -105.696
16.6556 | -105.7038
16.6593 | -105.7099
16.663 | -105.7193
16.6667 | -105.7298
16.6704 | -105.737
16.6741 | -105.7437
16.6778 | -105.7501
16.6815 | -105.7596
16.6852 | -105.7701
16.6889 | -105.7808
16.6926 | -105.7886
16.6963 | -105.7949
16.7 | -105.8035
16.7037 | -105.8143
16.7074 | -105.823
16.7111 | -105.8316
16.7148 | -105.841
16.7185 | -105.852
16.7222 | -105.864
16.7259 | -105.8717
16.7296 | -105.8788
16.7333 | -105.8865
16.737 | -105.899
16.7407 | -105.9079
16.7444 | -105.9148
16.7481 | -105.9225
16.7519 | -105.9286
16.7556 | -105.9327
16.7593 | -105.939
16.763 | -105.9446
16.7667 | -105.9481
16.7704 | -105.9505
16.7741 | -105.9527
16.7778 | -105.9522
16.7815 | -105.9524
16.7852 | -105.9514
16.7889 | -105.9496
16.7926 | -105.9465
16.7963 | -105.9422
16.8 | -105.9386
16.8037 | -105.9363
16.8074 | -105.9307
16.8111 | -105.9246
16.8148 | -105.9182
16.8185 | -105.9115
16.8222 | -105.9018
16.8259 | -105.8948
16.8296 | -105.8869
16.8333 | -105.8772
16.837 | -105.8648
16.8407 | -105.8521
16.8444 | -105.841
16.8481 | -105.8297
16.8519 | -105.8182
16.8556 | -105.8095
16.8593 | -105.8028
16.863 | -105.7988
16.8667 | -105.7924
16.8704 | -105.7861
16.8741 | -105.7808
16.8778 | -105.7736
16.8815 | -105.7667
16.8852 | -105.7612
16.8889 | -105.7586
16.8926 | -105.7559
16.8963 | -105.7531
16.9 | -105.7522
16.9037 | -105.7526
16.9074 | -105.7532
16.9111 | -105.7519
16.9148 | -105.751
16.9185 | -105.7521
16.9222 | -105.7515
16.9259 | -105.7508
16.9296 | -105.7514
16.9333 | -105.7528
16.937 | -105.7564
16.9407 | -105.7591
16.9444 | -105.7618
16.9481 | -105.7648
16.9519 | -105.7677
16.9556 | -105.7703
16.9593 | -105.7741
16.963 | -105.7775
16.9667 | -105.7832
16.9704 | -105.7907
16.9741 | -105.7978
16.9778 | -105.802
16.9815 | -105.8077
16.9852 | -105.8133
16.9889 | -105.8171
16.9926 | -105.8212
16.9963 | -105.8284
17 | -105.8349
17.0037 | -105.8391
17.0074 | -105.8444
17.0111 | -105.8533
17.0148 | -105.863
17.0185 | -105.8719
17.0222 | -105.8803
17.0259 | -105.8874
17.0296 | -105.8956
17.0333 | -105.9042
17.037 | -105.9116
17.0407 | -105.9178
17.0444 | -105.9201
17.0481 | -105.9233
17.0519 | -105.9287
17.0556 | -105.9309
17.0593 | -105.9334
17.063 | -105.9357
17.0667 | -105.9349
17.0704 | -105.9326
17.0741 | -105.9313
17.0778 | -105.9305
17.0815 | -105.9275
17.0852 | -105.9203
17.0889 | -105.9121
17.0926 | -105.9023
17.0963 | -105.8856
17.1 | -105.8731
17.1037 | -105.8599
17.1074 | -105.8463
17.1111 | -105.8342
17.1148 | -105.821
17.1185 | -105.8077
17.1222 | -105.796
17.1259 | -105.7846
17.1296 | -105.773
17.1333 | -105.7601
17.137 | -105.748
17.1407 | -105.7368
17.1444 | -105.729
17.1481 | -105.7198
17.1519 | -105.7126
17.1556 | -105.7104
17.1593 | -105.7059
17.163 | -105.7037
17.1667 | -105.7024
17.1704 | -105.7022
17.1741 | -105.7019
17.1778 | -105.7031
17.1815 | -105.7062
17.1852 | -105.7106
17.1889 | -105.7142
17.1926 | -105.7196
17.1963 | -105.7259
17.2 | -105.7331
17.2037 | -105.7407
17.2074 | -105.7473
17.2111 | -105.7553
17.2148 | -105.762
17.2185 | -105.7686
17.2222 | -105.7764
17.2259 | -105.7818
17.2296 | -105.7869
17.2333 | -105.7943
17.237 | -105.8017
17.2407 | -105.8047
17.2444 | -105.8047
17.2481 | -105.8077
17.2519 | -105.8111
17.2556 | -105.8118
17.2593 | -105.8124
17.263 | -105.8118
17.2667 | -105.8099
17.2704 | -105.8101
17.2741 | -105.8088
17.2778 | -105.8085
17.2815 | -105.8082
17.2852 | -105.8089
17.2889 | -105.8124
17.2926 | -105.8143
17.2963 | -105.8147
17.3 | -105.8165
17.3037 | -105.8206
17.3074 | -105.8247
17.3111 | -105.8313
17.3148 | -105.8396
17.3185 | -105.8459
17.3222 | -105.8529
17.3259 | -105.8596
17.3296 | -105.8688
17.3333 | -105.8792
17.337 | -105.8902
17.3407 | -105.9034
17.3444 | -105.9187
17.3481 | -105.9341
17.3519 | -105.9483
17.3556 | -105.9631
17.3593 | -105.9763
17.363 | -105.9875
17.3667 | -105.9984
17.3704 | -106.0094
17.3741 | -106.0185
17.3778 | -106.0259
17.3815 | -106.0331
17.3852 | -106.0388
17.3889 | -106.0432
17.3926 | -106.0445
17.3963 | -106.0451
17.4 | -106.0402
17.4037 | -106.0327
17.4074 | -106.0256
17.4111 | -106.0144
17.4148 | -106.0008
17.4185 | -105.9864
17.4222 | -105.9699
17.4259 | -105.9479
17.4296 | -105.9269
17.4333 | -105.903
17.437 | -105.8788
17.4407 | -105.854
17.4444 | -105.8284
17.4481 | -105.8013
17.4519 | -105.7732
17.4556 | -105.7474
17.4593 | -105.7225
17.463 | -105.6976
17.4667 | -105.6729
17.4704 | -105.6498
17.4741 | -105.6291
17.4778 | -105.6095
17.4815 | -105.5879
17.4852 | -105.57
17.4889 | -105.5563
17.4926 | -105.5446
17.4963 | -105.5327
17.5 | -105.5209
17.5037 | -105.5115
17.5074 | -105.5044
17.5111 | -105.5011
17.5148 | -105.4981
17.5185 | -105.4985
17.5222 | -105.5027
17.5259 | -105.5086
17.5296 | -105.515
17.5333 | -105.5199
17.537 | -105.5296
17.5407 | -105.5426
17.5444 | -105.5567
17.5481 | -105.5732
17.5519 | -105.587
17.5556 | -105.6005
17.5593 | -105.6163
17.563 | -105.6327
17.5667 | -105.6507
17.5704 | -105.6683
17.5741 | -105.6864
17.5778 | -105.7069
17.5815 | -105.7262
17.5852 | -105.7431
17.5889 | -105.7563
17.5926 | -105.7673
17.5963 | -105.7768
17.6 | -105.7875
17.6037 | -105.796
17.6074 | -105.8028
17.6111 | -105.8102
17.6148 | -105.814
17.6185 | -105.8166
17.6222 | -105.8178
17.6259 | -105.8195
17.6296 | -105.8106
17.6333 | -105.8063
17.637 | -105.8017
17.6407 | -105.7958
17.6444 | -105.7889
17.6481 | -105.7824
17.6519 | -105.7763
17.6556 | -105.7661
17.6593 | -105.7558
17.663 | -105.7427
17.6667 | -105.7313
17.6704 | -105.7203
17.6741 | -105.7087
17.6778 | -105.701
17.6815 | -105.6891
17.6852 | -105.6783
17.6889 | -105.6685
17.6926 | -105.6576
17.6963 | -105.6431
17.7 | -105.6313
17.7037 | -105.6204
17.7074 | -105.6088
17.7111 | -105.5974
17.7148 | -105.587
17.7185 | -105.5787
17.7222 | -105.5695
17.7259 | -105.5595
17.7296 | -105.5503
17.7333 | -105.5415
17.737 | -105.5336
17.7407 | -105.5252
17.7444 | -105.5186
17.7481 | -105.5144
17.7519 | -105.5094
17.7556 | -105.506
17.7593 | -105.5036
17.763 | -105.5015
17.7667 | -105.501
17.7704 | -105.4988
17.7741 | -105.4964
17.7778 | -105.4972
17.7815 | -105.5009
17.7852 | -105.5065
17.7889 | -105.51
17.7926 | -105.5145
17.7963 | -105.5212
17.8 | -105.5279
17.8037 | -105.5376
17.8074 | -105.5505
17.8111 | -105.561
17.8148 | -105.5716
17.8185 | -105.5838
17.8222 | -105.5944
17.8259 | -105.6053
17.8296 | -105.6157
17.8333 | -105.6251
17.837 | -105.6347
17.8407 | -105.6459
17.8444 | -105.6574
17.8481 | -105.6686
17.8519 | -105.6781
17.8556 | -105.6894
17.8593 | -105.699
17.863 | -105.7062
17.8667 | -105.715
17.8704 | -105.7207
17.8741 | -105.7252
17.8778 | -105.7281
17.8815 | -105.7315
17.8852 | -105.7358
17.8889 | -105.7389
17.8926 | -105.7409
17.8963 | -105.7403
17.9 | -105.7343
17.9037 | -105.7272
17.9074 | -105.7214
17.9111 | -105.7176
17.9148 | -105.7109
17.9185 | -105.7025
17.9222 | -105.6921
17.9259 | -105.6782
17.9296 | -105.6635
17.9333 | -105.6467
17.937 | -105.6316
17.9407 | -105.6148
17.9444 | -105.5952
17.9481 | -105.5749
17.9519 | -105.5554
17.9556 | -105.538
17.9593 | -105.5199
17.963 | -105.501
17.9667 | -105.4819
17.9704 | -105.4623
17.9741 | -105.4432
17.9778 | -105.4276
17.9815 | -105.415
17.9852 | -105.4037
17.9889 | -105.3934
17.9926 | -105.3819
17.9963 | -105.3736
18 | -105.368
18.0037 | -105.3634
18.0074 | -105.3628
18.0111 | -105.3616
18.0148 | -105.3622
18.0185 | -105.3648
18.0222 | -105.3704
18.0259 | -105.374
18.0296 | -105.3804
18.0333 | -105.3872
18.037 | -105.3938
18.0407 | -105.4025
18.0444 | -105.4111
18.0481 | -105.4228
18.0519 | -105.4341
18.0556 | -105.4445
18.0593 | -105.4552
18.063 | -105.4641
18.0667 | -105.4741
18.0704 | -105.4843
18.0741 | -105.4941
18.0778 | -105.5069
18.0815 | -105.5175
18.0852 | -105.5257
18.0889 | -105.5352
18.0926 | -105.5421
18.0963 | -105.5465
18.1 | -105.5506
18.1037 | -105.5516
18.1074 | -105.5521
18.1111 | -105.5503
18.1148 | -105.5494
18.1185 | -105.5471
18.1222 | -105.5409
18.1259 | -105.5352
18.1296 | -105.5312
18.1333 | -105.5258
18.137 | -105.5189
18.1407 | -105.5115
18.1444 | -105.5033
18.1481 | -105.4947
18.1519 | -105.4863
18.1556 | -105.4821
18.1593 | -105.4774
18.163 | -105.4712
18.1667 | -105.4666
18.1704 | -105.4588
18.1741 | -105.4533
18.1778 | -105.4444
18.1815 | -105.4327
18.1852 | -105.4224
18.1889 | -105.4143
18.1926 | -105.4061
18.1963 | -105.3967
18.2 | -105.3915
18.2037 | -105.387
18.2074 | -105.3819
18.2111 | -105.3774
18.2148 | -105.368
18.2185 | -105.3576
18.2222 | -105.35
18.2259 | -105.342
18.2296 | -105.3362
18.2333 | -105.3302
18.237 | -105.3211
18.2407 | -105.312
18.2444 | -105.3044
18.2481 | -105.2973
18.2519 | -105.2911
18.2556 | -105.2852
18.2593 | -105.2785
18.263 | -105.2712
18.2667 | -105.2635
18.2704 | -105.2568
18.2741 | -105.249
18.2778 | -105.2422
18.2815 | -105.2371
18.2852 | -105.2348
18.2889 | -105.2317
18.2926 | -105.2262
18.2963 | -105.2196
18.3 | -105.2119
18.3037 | -105.2092
18.3074 | -105.2061
18.3111 | -105.2019
18.3148 | -105.1991
18.3185 | -105.1983
18.3222 | -105.2002
18.3259 | -105.2009
18.3296 | -105.2014
18.3333 | -105.1998
18.337 | -105.1974
18.3407 | -105.1953
18.3444 | -105.1942
18.3481 | -105.1929
18.3519 | -105.1934
18.3556 | -105.195
18.3593 | -105.1965
18.363 | -105.1984
18.3667 | -105.199
18.3704 | -105.1972
18.3741 | -105.1941
18.3778 | -105.1938
18.3815 | -105.1942
18.3852 | -105.1933
18.3889 | -105.1916
18.3926 | -105.1898
18.3963 | -105.1889
18.4 | -105.1892
18.4037 | -105.1897
18.4074 | -105.189
18.4111 | -105.1903
18.4148 | -105.1897
18.4185 | -105.1906
18.4222 | -105.19
18.4259 | -105.1919
18.4296 | -105.1981
18.4333 | -105.2034
18.437 | -105.2081
18.4407 | -105.2185
18.4444 | -105.2286
18.4481 | -105.2395
18.4519 | -105.2478
18.4556 | -105.2563
18.4593 | -105.2656
18.463 | -105.2722
18.4667 | -105.2802
18.4704 | -105.2896
18.4741 | -105.3009
18.4778 | -105.3138
18.4815 | -105.3283
18.4852 | -105.3408
18.4889 | -105.3538
18.4926 | -105.3663
18.4963 | -105.3792
18.5 | -105.3921
18.5037 | -105.4033
18.5074 | -105.4136
18.5111 | -105.4251
18.5148 | -105.4355
18.5185 | -105.4449
18.5222 | -105.4538
18.5259 | -105.4621
18.5296 | -105.4705
18.5333 | -105.4777
18.537 | -105.4838
18.5407 | -105.4884
18.5444 | -105.4932
18.5481 | -105.4955
18.5519 | -105.496
18.5556 | -105.5001
18.5593 | -105.5061
18.563 | -105.5078
18.5667 | -105.5053
18.5704 | -105.5027
18.5741 | -105.5017
18.5778 | -105.5002
18.5815 | -105.4943
18.5852 | -105.4866
18.5889 | -105.4797
18.5926 | -105.4731
18.5963 | -105.4698
18.6 | -105.4661
18.6037 | -105.4613
18.6074 | -105.4563
18.6111 | -105.4505
18.6148 | -105.4429
18.6185 | -105.4358
18.6222 | -105.4277
18.6259 | -105.4216
18.6296 | -105.4173
18.6333 | -105.4145
18.637 | -105.4112
18.6407 | -105.4039
18.6444 | -105.3981
18.6481 | -105.3937
18.6519 | -105.389
18.6556 | -105.3864
18.6593 | -105.3835
18.663 | -105.3788
18.6667 | -105.3777
18.6704 | -105.3763
18.6741 | -105.3743
18.6778 | -105.3705
18.6815 | -105.3627
18.6852 | -105.3587
18.6889 | -105.3528
18.6926 | -105.3466
18.6963 | -105.3415
18.7 | -105.3288
18.7037 | -105.3191
18.7074 | -105.3098
18.7111 | -105.3033
18.7148 | -105.294
18.7185 | -105.2838
18.7222 | -105.2713
18.7259 | -105.2563
18.7296 | -105.2421
18.7333 | -105.2263
18.737 | -105.2117
18.7407 | -105.1972
18.7444 | -105.1802
18.7481 | -105.1608
18.7519 | -105.1438
18.7556 | -105.1266
18.7593 | -105.1078
18.763 | -105.0873
18.7667 | -105.0678
18.7704 | -105.0518
18.7741 | -105.0367
18.7778 | -105.0206
18.7815 | -105.0062
18.7852 | -104.9893
18.7889 | -104.9737
18.7926 | -104.9607
18.7963 | -104.9492
18.8 | -104.9372
18.8037 | -104.9278
18.8074 | -104.9192
18.8111 | -104.9102
18.8148 | -104.9053
18.8185 | -104.9027
18.8222 | -104.8984
18.8259 | -104.8969
18.8296 | -104.8988
18.8333 | -104.8984
18.837 | -104.8987
18.8407 | -104.9009
18.8444 | -104.905
18.8481 | -104.9131
18.8519 | -104.9235
18.8556 | -104.9344
18.8593 | -104.9467
18.863 | -104.9582
18.8667 | -104.9697
18.8704 | -104.985
18.8741 | -105.0007
18.8778 | -105.017
18.8815 | -105.0359
18.8852 | -105.0535
18.8889 | -105.0719
18.8926 | -105.0915
18.8963 | -105.1097
18.9 | -105.1292
18.9037 | -105.1492
18.9074 | -105.1701
18.9111 | -105.1918
18.9148 | -105.2129
18.9185 | -105.234
18.9222 | -105.2532
18.9259 | -105.2726
18.9296 | -105.2922
18.9333 | -105.3094
18.937 | -105.3249
18.9407 | -105.339
18.9444 | -105.3533
18.9481 | -105.3686
18.9519 | -105.3813
18.9556 | -105.3925
18.9593 | -105.3995
18.963 | -105.4011
18.9667 | -105.4066
18.9704 | -105.4115
18.9741 | -105.4138
18.9778 | -105.4123
18.9815 | -105.4085
18.9852 | -105.4042
18.9889 | -105.3971
18.9926 | -105.3894
18.9963 | -105.3796
19 | -105.3642
19.0037 | -105.3482
19.0074 | -105.3292
19.0111 | -105.309
19.0148 | -105.2905
19.0185 | -105.2729
19.0222 | -105.249
19.0259 | -105.2258
19.0296 | -105.2045
19.0333 | -105.1826
19.037 | -105.1603
19.0407 | -105.1375
19.0444 | -105.1144
19.0481 | -105.0912
19.0519 | -105.0673
19.0556 | -105.0475
19.0593 | -105.0269
19.063 | -105.0081
19.0667 | -104.9907
19.0704 | -104.9739
19.0741 | -104.9591
19.0778 | -104.9439
19.0815 | -104.931
19.0852 | -104.9192
19.0889 | -104.9085
19.0926 | -104.9011
19.0963 | -104.8941
19.1 | -104.8897
19.1037 | -104.8882
19.1074 | -104.8866
19.1111 | -104.8853
19.1148 | -104.8856
19.1185 | -104.8855
19.1222 | -104.8844
19.1259 | -104.8858
19.1296 | -104.8873
19.1333 | -104.8909
19.137 | -104.895
19.1407 | -104.8962
19.1444 | -104.8973
19.1481 | -104.8994
19.1519 | -104.8995
19.1556 | -104.8997
19.1593 | -104.8993
19.163 | -104.8972
19.1667 | -104.8963
19.1704 | -104.8934
19.1741 | -104.8895
19.1778 | -104.8829
19.1815 | -104.8746
19.1852 | -104.8646
19.1889 | -104.8534
19.1926 | -104.8414
19.1963 | -104.8285
19.2 | -104.8144
19.2037 | -104.8003
19.2074 | -104.7852
19.2111 | -104.7678
19.2148 | -104.7505
19.2185 | -104.7315
19.2222 | -104.7115
19.2259 | -104.6881
19.2296 | -104.6641
19.2333 | -104.6414
19.237 | -104.6233
19.2407 | -104.6019
19.2444 | -104.5766
19.2481 | -104.5536
19.2519 | -104.5323
19.2556 | -104.5103
19.2593 | -104.4873
19.263 | -104.466
19.2667 | -104.445
19.2704 | -104.4256
19.2741 | -104.4063
19.2778 | -104.3857
19.2815 | -104.3664
19.2852 | -104.347
19.2889 | -104.3283
19.2926 | -104.3103
19.2963 | -104.2932
19.3 | -104.2749
19.3037 | -104.2559
19.3074 | -104.2401
19.3111 | -104.2253
19.3148 | -104.2084
19.3185 | -104.1934
19.3222 | -104.1794
19.3259 | -104.1645
19.3296 | -104.1475
19.3333 | -104.1305
19.337 | -104.1142
19.3407 | -104.0987
19.3444 | -104.0849
19.3481 | -104.0698
19.3519 | -104.0552
19.3556 | -104.0391
19.3593 | -104.023
19.363 | -104.0088
19.3667 | -103.9946
19.3704 | -103.9814
19.3741 | -103.9685
19.3778 | -103.9549
19.3815 | -103.942
19.3852 | -103.9331
19.3889 | -103.9228
19.3926 | -103.913
19.3963 | -103.9051
19.4 | -103.8975
19.4037 | -103.8898
19.4074 | -103.8823
19.4111 | -103.8767
19.4148 | -103.8732
19.4185 | -103.8696
19.4222 | -103.8658
19.4259 | -103.8611
19.4296 | -103.8589
19.4333 | -103.857
19.437 | -103.8546
19.4407 | -103.853
19.4444 | -103.8534
19.4481 | -103.8535
19.4519 | -103.8561
19.4556 | -103.8576
19.4593 | -103.8593
19.463 | -103.862
19.4667 | -103.8659
19.4704 | -103.8706
19.4741 | -103.8757
19.4778 | -103.8786
19.4815 | -103.8859
19.4852 | -103.8934
19.4889 | -103.8996
19.4926 | -103.9061
19.4963 | -103.9136
19.5 | -103.9231
19.5037 | -103.9336
19.5074 | -103.942
19.5111 | -103.9502
19.5148 | -103.9566
19.5185 | -103.9636
19.5222 | -103.9737
19.5259 | -103.9823
19.5296 | -103.9889
19.5333 | -103.9943
19.537 | -104.0005
19.5407 | -104.0076
19.5444 | -104.016
19.5481 | -104.0234
19.5519 | -104.0313
19.5556 | -104.0386
19.5593 | -104.0442
19.563 | -104.047
19.5667 | -104.0492
19.5704 | -104.0515
19.5741 | -104.0559
19.5778 | -104.0631
19.5815 | -104.07
19.5852 | -104.0745
19.5889 | -104.0786
19.5926 | -104.0848
19.5963 | -104.0902
19.6 | -104.098
19.6037 | -104.1082
19.6074 | -104.1203
19.6111 | -104.1336
19.6148 | -104.1452
19.6185 | -104.1579
19.6222 | -104.1716
19.6259 | -104.1858
19.6296 | -104.2026
19.6333 | -104.2192
19.637 | -104.2365
19.6407 | -104.2544
19.6444 | -104.2742
19.6481 | -104.2917
19.6519 | -104.3095
19.6556 | -104.331
19.6593 | -104.3534
19.663 | -104.3729
19.6667 | -104.3942
19.6704 | -104.4147
19.6741 | -104.4348
19.6778 | -104.4539
19.6815 | -104.4726
19.6852 | -104.4922
19.6889 | -104.5113
19.6926 | -104.5301
19.6963 | -104.5495
19.7 | -104.5685
19.7037 | -104.5895
19.7074 | -104.605
19.7111 | -104.6185
19.7148 | -104.6326
19.7185 | -104.6461
19.7222 | -104.6585
19.7259 | -104.6707
19.7296 | -104.6821
19.7333 | -104.6932
19.737 | -104.7016
19.7407 | -104.7078
19.7444 | -104.7117
19.7481 | -104.7133
19.7519 | -104.7143
19.7556 | -104.7182
19.7593 | -104.7175
19.763 | -104.7165
19.7667 | -104.7042
19.7704 | -104.7024
19.7741 | -104.6989
19.7778 | -104.6926
19.7815 | -104.6851
19.7852 | -104.6774
19.7889 | -104.6686
19.7926 | -104.6577
19.7963 | -104.6483
19.8 | -104.6382
19.8037 | -104.6281
19.8074 | -104.6169
19.8111 | -104.6049
19.8148 | -104.5941
19.8185 | -104.5818
19.8222 | -104.5681
19.8259 | -104.551
19.8296 | -104.5339
19.8333 | -104.5207
19.837 | -104.508
19.8407 | -104.4957
19.8444 | -104.486
19.8481 | -104.476
19.8519 | -104.466
19.8556 | -104.4565
19.8593 | -104.4465
19.863 | -104.4374
19.8667 | -104.4282
19.8704 | -104.4189
19.8741 | -104.4084
19.8778 | -104.3991
19.8815 | -104.3914
19.8852 | -104.3844
19.8889 | -104.3777
19.8926 | -104.3691
19.8963 | -104.3607
19.9 | -104.3533
19.9037 | -104.3453
19.9074 | -104.3374
19.9111 | -104.3312
19.9148 | -104.3261
19.9185 | -104.3212
19.9222 | -104.3152
19.9259 | -104.3058
19.9296 | -104.2997
19.9333 | -104.2943
19.937 | -104.2857
19.9407 | -104.2775
19.9444 | -104.2719
19.9481 | -104.2655
19.9519 | -104.2609
19.9556 | -104.258
19.9593 | -104.2559
19.963 | -104.2543
19.9667 | -104.2494
19.9704 | -104.2448
19.9741 | -104.2399
19.9778 | -104.2336
19.9815 | -104.2282
19.9852 | -104.2241
19.9889 | -104.2222
19.9926 | -104.2185
19.9963 | -104.2165
20 | -104.2129
20.0037 | -104.2101
20.0074 | -104.2089
20.0111 | -104.2078
20.0148 | -104.2075
20.0185 | -104.2087
20.0222 | -104.2092
20.0259 | -104.2099
20.0296 | -104.2139
20.0333 | -104.2184
20.037 | -104.2198
20.0407 | -104.2219
20.0444 | -104.2258
20.0481 | -104.2287
20.0519 | -104.233
20.0556 | -104.236
20.0593 | -104.2402
20.063 | -104.2438
20.0667 | -104.2471
20.0704 | -104.2503
20.0741 | -104.2528
20.0778 | -104.2559
20.0815 | -104.259
20.0852 | -104.2612
20.0889 | -104.2608
20.0926 | -104.2594
20.0963 | -104.2586
20.1 | -104.257
20.1037 | -104.2537
20.1074 | -104.2492
20.1111 | -104.2451
20.1148 | -104.2405
20.1185 | -104.2333
20.1222 | -104.2253
20.1259 | -104.2142
20.1296 | -104.2031
20.1333 | -104.1914
20.137 | -104.1792
20.1407 | -104.1675
20.1444 | -104.1558
20.1481 | -104.1449
20.1519 | -104.1341
20.1556 | -104.1243
20.1593 | -104.1155
20.163 | -104.1039
20.1667 | -104.0929
20.1704 | -104.0821
20.1741 | -104.0726
20.1778 | -104.0651
20.1815 | -104.0557
20.1852 | -104.0438
20.1889 | -104.0359
20.1926 | -104.0277
20.1963 | -104.0154
20.2 | -104.0036
20.2037 | -103.9942
20.2074 | -103.9856
20.2111 | -103.9806
20.2148 | -103.9748
20.2185 | -103.9676
20.2222 | -103.9588
20.2259 | -103.9468
20.2296 | -103.9355
20.2333 | -103.9262
20.237 | -103.9172
20.2407 | -103.9067
20.2444 | -103.897
20.2481 | -103.8887
20.2519 | -103.8843
20.2556 | -103.8801
20.2593 | -103.8754
20.263 | -103.8721
20.2667 | -103.868
20.2704 | -103.8618
20.2741 | -103.857
20.2778 | -103.8542
20.2815 | -103.8492
20.2852 | -103.8441
20.2889 | -103.8367
20.2926 | -103.8294
20.2963 | -103.8225
20.3 | -103.8159
20.3037 | -103.8081
20.3074 | -103.8028
20.3111 | -103.7957
20.3148 | -103.7883
20.3185 | -103.7801
20.3222 | -103.7719
20.3259 | -103.7641
20.3296 | -103.7555
20.3333 | -103.7464
20.337 | -103.7374
20.3407 | -103.7281
20.3444 | -103.7188
20.3481 | -103.7092
20.3519 | -103.7015
20.3556 | -103.6944
20.3593 | -103.6871
20.363 | -103.6785
20.3667 | -103.6707
20.3704 | -103.663
20.3741 | -103.6558
20.3778 | -103.6503
20.3815 | -103.6448
20.3852 | -103.6384
20.3889 | -103.6314
20.3926 | -103.6276
20.3963 | -103.6239
20.4 | -103.6195
20.4037 | -103.6159
20.4074 | -103.6114
20.4111 | -103.607
20.4148 | -103.6051
20.4185 | -103.6026
20.4222 | -103.6002
20.4259 | -103.5981
20.4296 | -103.5954
20.4333 | -103.5917
20.437 | -103.5873
20.4407 | -103.5844
20.4444 | -103.5803
20.4481 | -103.5764
20.4519 | -103.5744
20.4556 | -103.5693
20.4593 | -103.5649
20.463 | -103.5581
20.4667 | -103.5514
20.4704 | -103.55
20.4741 | -103.5491
20.4778 | -103.5449
20.4815 | -103.5425
20.4852 | -103.5413
20.4889 | -103.5408
20.4926 | -103.5382
20.4963 | -103.5368
20.5 | -103.5352
20.5037 | -103.5339
20.5074 | -103.5341
20.5111 | -103.5385
20.5148 | -103.5448
20.5185 | -103.5491
20.5222 | -103.5527
20.5259 | -103.5559
20.5296 | -103.5608
20.5333 | -103.5656
20.537 | -103.569
20.5407 | -103.5691
20.5444 | -103.5707
20.5481 | -103.5752
20.5519 | -103.5797
20.5556 | -103.5847
20.5593 | -103.59
20.563 | -103.5939
20.5667 | -103.5988
20.5704 | -103.6035
20.5741 | -103.6065
20.5778 | -103.6118
20.5815 | -103.6177
20.5852 | -103.6218
20.5889 | -103.6237
20.5926 | -103.6279
20.5963 | -103.6292
20.6 | -103.6298
20.6037 | -103.6318
20.6074 | -103.6333
20.6111 | -103.6336
20.6148 | -103.634
20.6185 | -103.6364
20.6222 | -103.6384
20.6259 | -103.6395
20.6296 | -103.6421
20.6333 | -103.644
20.637 | -103.6451
20.6407 | -103.6469
20.6444 | -103.6491
20.6481 | -103.6502
20.6519 | -103.6505
20.6556 | -103.6518
20.6593 | -103.6531
20.663 | -103.6557
20.6667 | -103.6579
20.6704 | -103.66
20.6741 | -103.6644
20.6778 | -103.6701
20.6815 | -103.6767
20.6852 | -103.6822
20.6889 | -103.6863
20.6926 | -103.691
20.6963 | -103.6957
20.7 | -103.7021
20.7037 | -103.7077
20.7074 | -103.7112
20.7111 | -103.714
20.7148 | -103.7185
20.7185 | -103.7234
20.7222 | -103.7258
20.7259 | -103.7296
20.7296 | -103.7346
20.7333 | -103.7405
20.737 | -103.7464
20.7407 | -103.7499
20.7444 | -103.7543
20.7481 | -103.759
20.7519 | -103.7633
20.7556 | -103.7682
20.7593 | -103.7715
20.763 | -103.771
20.7667 | -103.7723
20.7704 | -103.7747
20.7741 | -103.7759
20.7778 | -103.7769
20.7815 | -103.7776
20.7852 | -103.7777
20.7889 | -103.7754
20.7926 | -103.7745
20.7963 | -103.773
20.8 | -103.7699
20.8037 | -103.7671
20.8074 | -103.7631
20.8111 | -103.7592
20.8148 | -103.7563
20.8185 | -103.7553
20.8222 | -103.7526
20.8259 | -103.7509
20.8296 | -103.7503
20.8333 | -103.7514
20.837 | -103.757
20.8407 | -103.7583
20.8444 | -103.7588
20.8481 | -103.7622
20.8519 | -103.7666
20.8556 | -103.7715
20.8593 | -103.7774
20.863 | -103.7813
20.8667 | -103.7848
20.8704 | -103.7919
20.8741 | -103.8007
20.8778 | -103.8082
20.8815 | -103.8181
20.8852 | -103.8307
20.8889 | -103.8421
20.8926 | -103.8562
20.8963 | -103.8712
20.9 | -103.8857
20.9037 | -103.9005
20.9074 | -103.915
20.9111 | -103.9301
20.9148 | -103.9448
20.9185 | -103.9597
20.9222 | -103.9741
20.9259 | -103.9876
20.9296 | -104.0014
20.9333 | -104.0138
20.937 | -104.0262
20.9407 | -104.0383
20.9444 | -104.0491
20.9481 | -104.0613
20.9519 | -104.0729
20.9556 | -104.0821
20.9593 | -104.0895
20.963 | -104.0966
20.9667 | -104.1017
20.9704 | -104.1054
20.9741 | -104.1079
20.9778 | -104.1099
20.9815 | -104.1113
20.9852 | -104.1131
20.9889 | -104.1141
20.9926 | -104.1148
20.9963 | -104.1139
21 | -104.1154
21.0037 | -104.1166
21.0074 | -104.1162
21.0111 | -104.1126
21.0148 | -104.1073
21.0185 | -104.1051
21.0222 | -104.1019
21.0259 | -104.0968
21.0296 | -104.0913
21.0333 | -104.0855
21.037 | -104.0794
21.0407 | -104.0714
21.0444 | -104.0642
21.0481 | -104.0547
21.0519 | -104.0481
21.0556 | -104.0421
21.0593 | -104.034
21.063 | -104.0232
21.0667 | -104.0099
21.0704 | -103.9971
21.0741 | -103.9845
21.0778 | -103.9707
21.0815 | -103.957
21.0852 | -103.9457
21.0889 | -103.9328
21.0926 | -103.9195
21.0963 | -103.9079
21.1 | -103.8946
21.1037 | -103.8838
21.1074 | -103.8727
21.1111 | -103.8623
21.1148 | -103.8491
21.1185 | -103.8366
21.1222 | -103.8235
21.1259 | -103.8115
21.1296 | -103.8005
21.1333 | -103.7887
21.137 | -103.7787
21.1407 | -103.7676
21.1444 | -103.7566
21.1481 | -103.7454
21.1519 | -103.7352
21.1556 | -103.7263
21.1593 | -103.7173
21.163 | -103.7114
21.1667 | -103.7061
21.1704 | -103.7007
21.1741 | -103.6946
21.1778 | -103.6884
21.1815 | -103.6836
21.1852 | -103.6788
21.1889 | -103.6741
21.1926 | -103.6697
21.1963 | -103.6647
21.2 | -103.66
21.2037 | -103.6559
21.2074 | -103.6528
21.2111 | -103.6494
21.2148 | -103.6459
21.2185 | -103.644
21.2222 | -103.6444
21.2259 | -103.6449
21.2296 | -103.6437
21.2333 | -103.643
21.237 | -103.6418
21.2407 | -103.6394
21.2444 | -103.6391
21.2481 | -103.6398
21.2519 | -103.6382
21.2556 | -103.6374
21.2593 | -103.6357
21.263 | -103.6326
21.2667 | -103.63
21.2704 | -103.6271
21.2741 | -103.6233
21.2778 | -103.618
21.2815 | -103.6114
21.2852 | -103.6039
21.2889 | -103.5973
21.2926 | -103.5896
21.2963 | -103.5827
21.3 | -103.5739
21.3037 | -103.5648
21.3074 | -103.5571
21.3111 | -103.5487
21.3148 | -103.54
21.3185 | -103.5287
21.3222 | -103.5176
21.3259 | -103.507
21.3296 | -103.495
21.3333 | -103.4858
21.337 | -103.479
21.3407 | -103.4703
21.3444 | -103.4589
21.3481 | -103.4515
21.3519 | -103.4407
21.3556 | -103.4287
21.3593 | -103.419
21.363 | -103.4101
21.3667 | -103.4049
21.3704 | -103.3972
21.3741 | -103.3908
21.3778 | -103.3857
21.3815 | -103.38
21.3852 | -103.3747
21.3889 | -103.3683
21.3926 | -103.364
21.3963 | -103.3608
21.4 | -103.3572
21.4037 | -103.3545
21.4074 | -103.3506
21.4111 | -103.3478
21.4148 | -103.3452
21.4185 | -103.3453
21.4222 | -103.3456
21.4259 | -103.3458
21.4296 | -103.3462
21.4333 | -103.3476
21.437 | -103.3501
21.4407 | -103.3515
21.4444 | -103.3533
21.4481 | -103.3545
21.4519 | -103.3555
21.4556 | -103.3569
21.4593 | -103.3578
21.463 | -103.3612
21.4667 | -103.3629
21.4704 | -103.3644
21.4741 | -103.3652
21.4778 | -103.3653
21.4815 | -103.3666
21.4852 | -103.3663
21.4889 | -103.3686
21.4926 | -103.3711
21.4963 | -103.3726
21.5 | -103.3769
21.5037 | -103.3807
21.5074 | -103.3825
21.5111 | -103.384
21.5148 | -103.3854
21.5185 | -103.3879
21.5222 | -103.3909
21.5259 | -103.3928
21.5296 | -103.3948
21.5333 | -103.398
21.537 | -103.3998
21.5407 | -103.3999
21.5444 | -103.3995
21.5481 | -103.3995
21.5519 | -103.3978
21.5556 | -103.3976
21.5593 | -103.3974
21.563 | -103.3968
21.5667 | -103.3972
21.5704 | -103.3963
21.5741 | -103.3941
21.5778 | -103.3907
21.5815 | -103.3874
21.5852 | -103.3857
21.5889 | -103.3841
21.5926 | -103.38
21.5963 | -103.3745
21.6 | -103.3699
21.6037 | -103.3656
21.6074 | -103.3615
21.6111 | -103.3569
21.6148 | -103.3515
21.6185 | -103.3466
21.6222 | -103.3403
21.6259 | -103.3346
21.6296 | -103.3301
21.6333 | -103.3265
21.637 | -103.3214
21.6407 | -103.3145
21.6444 | -103.3095
21.6481 | -103.3076
21.6519 | -103.3029
21.6556 | -103.2979
21.6593 | -103.2957
21.663 | -103.2907
21.6667 | -103.2861
21.6704 | -103.282
21.6741 | -103.2757
21.6778 | -103.2711
21.6815 | -103.2655
21.6852 | -103.2582
21.6889 | -103.2507
21.6926 | -103.2435
21.6963 | -103.2354
21.7 | -103.2257
21.7037 | -103.2147
21.7074 | -103.2015
21.7111 | -103.1903
21.7148 | -103.182
21.7185 | -103.1724
21.7222 | -103.1616
21.7259 | -103.1506
21.7296 | -103.14
21.7333 | -103.1301
21.737 | -103.1195
21.7407 | -103.1078
21.7444 | -103.0964
21.7481 | -103.085
21.7519 | -103.0731
21.7556 | -103.0627
21.7593 | -103.0537
21.763 | -103.0427
21.7667 | -103.0317
21.7704 | -103.0203
21.7741 | -103.0101
21.7778 | -103.0013
21.7815 | -102.9921
21.7852 | -102.982
21.7889 | -102.9732
21.7926 | -102.9674
21.7963 | -102.9594
21.8 | -102.9512
21.8037 | -102.9433
21.8074 | -102.9352
21.8111 | -102.9298
21.8148 | -102.9248
21.8185 | -102.9193
21.8222 | -102.9159
21.8259 | -102.9127
21.8296 | -102.9085
21.8333 | -102.9043
21.837 | -102.902
21.8407 | -102.902
21.8444 | -102.9035
21.8481 | -102.9056
21.8519 | -102.9075
21.8556 | -102.9097
21.8593 | -102.9091
21.863 | -102.9113
21.8667 | -102.9139
21.8704 | -102.9167
21.8741 | -102.9201
21.8778 | -102.9243
21.8815 | -102.9302
21.8852 | -102.9381
21.8889 | -102.9451
21.8926 | -102.9498
21.8963 | -102.9569
21.9 | -102.9659
21.9037 | -102.9737
21.9074 | -102.9824
21.9111 | -102.9893
21.9148 | -102.9967
21.9185 | -103.004
21.9222 | -103.0111
21.9259 | -103.0198
21.9296 | -103.0296
21.9333 | -103.0408
21.937 | -103.0491
21.9407 | -103.056
21.9444 | -103.0614
21.9481 | -103.067
21.9519 | -103.0739
21.9556 | -103.0811
21.9593 | -103.0867
21.963 | -103.0901
21.9667 | -103.0945
21.9704 | -103.0984
21.9741 | -103.1036
21.9778 | -103.1053
21.9815 | -103.1048
21.9852 | -103.1074
21.9889 | -103.1072
21.9926 | -103.1063
21.9963 | -103.1055
22 | -103.1028
22.0037 | -103.0976
22.0074 | -103.0925
22.0111 | -103.087
22.0148 | -103.0817
22.0185 | -103.0764
22.0222 | -103.0649
22.0259 | -103.0513
22.0296 | -103.0381
22.0333 | -103.0254
22.037 | -103.0087
22.0407 | -102.993
22.0444 | -102.9783
22.0481 | -102.9596
22.0519 | -102.9405
22.0556 | -102.9219
22.0593 | -102.9045
22.063 | -102.888
22.0667 | -102.8701
22.0704 | -102.8515
22.0741 | -102.8348
22.0778 | -102.815
22.0815 | -102.7925
22.0852 | -102.7717
22.0889 | -102.7527
22.0926 | -102.7359
22.0963 | -102.7199
22.1 | -102.7018
22.1037 | -102.6821
22.1074 | -102.665
22.1111 | -102.6495
22.1148 | -102.6348
22.1185 | -102.6172
22.1222 | -102.598
22.1259 | -102.5817
22.1296 | -102.5661
22.1333 | -102.5493
22.137 | -102.5332
22.1407 | -102.5175
22.1444 | -102.5033
22.1481 | -102.4887
22.1519 | -102.4737
22.1556 | -102.4599
22.1593 | -102.4477
22.163 | -102.4358
22.1667 | -102.4219
22.1704 | -102.4109
22.1741 | -102.4004
22.1778 | -102.3905
22.1815 | -102.3808
22.1852 | -102.3717
22.1889 | -102.3635
22.1926 | -102.3536
22.1963 | -102.3439
22.2 | -102.337
22.2037 | -102.3325
22.2074 | -102.3273
22.2111 | -102.3218
22.2148 | -102.3197
22.2185 | -102.3209
22.2222 | -102.3213
22.2259 | -102.3222
22.2296 | -102.3241
22.2333 | -102.3261
22.237 | -102.3294
22.2407 | -102.3343
22.2444 | -102.3383
22.2481 | -102.343
22.2519 | -102.3489
22.2556 | -102.3547
22.2593 | -102.3613
22.263 | -102.3684
22.2667 | -102.3766
22.2704 | -102.3853
22.2741 | -102.3927
22.2778 | -102.4024
22.2815 | -102.4117
22.2852 | -102.419
22.2889 | -102.4262
22.2926 | -102.4364
22.2963 | -102.4452
22.3 | -102.4522
22.3037 | -102.461
22.3074 | -102.4685
22.3111 | -102.4741
22.3148 | -102.4805
22.3185 | -102.4883
22.3222 | -102.496
22.3259 | -102.5019
22.3296 | -102.5076
22.3333 | -102.513
22.337 | -102.521
22.3407 | -102.5266
22.3444 | -102.5307
22.3481 | -102.5327
22.3519 | -102.5359
22.3556 | -102.5384
22.3593 | -102.5384
22.363 | -102.5395
22.3667 | -102.5394
22.3704 | -102.5373
22.3741 | -102.5346
22.3778 | -102.5326
22.3815 | -102.5292
22.3852 | -102.5258
22.3889 | -102.524
22.3926 | -102.5244
22.3963 | -102.5256
22.4 | -102.5247
22.4037 | -102.5207
22.4074 | -102.514
22.4111 | -102.5088
22.4148 | -102.5032
22.4185 | -102.4975
22.4222 | -102.4928
22.4259 | -102.49
22.4296 | -102.4878
22.4333 | -102.4855
22.437 | -102.4834
22.4407 | -102.4806
22.4444 | -102.4757
22.4481 | -102.4701
22.4519 | -102.4638
22.4556 | -102.4577
22.4593 | -102.4513
22.463 | -102.4439
22.4667 | -102.4375
22.4704 | -102.431
22.4741 | -102.4228
22.4778 | -102.4152
22.4815 | -102.409
22.4852 | -102.4041
22.4889 | -102.3987
22.4926 | -102.3887
22.4963 | -102.378
22.5 | -102.3684
22.5037 | -102.3583
22.5074 | -102.3475
22.5111 | -102.3369
22.5148 | -102.3255
22.5185 | -102.3126
22.5222 | -102.2985
22.5259 | -102.2836
22.5296 | -102.2683
22.5333 | -102.2533
22.537 | -102.2365
22.5407 | -102.2171
22.5444 | -102.1988
22.5481 | -102.1806
22.5519 | -102.1636
22.5556 | -102.1474
22.5593 | -102.1289
22.563 | -102.1104
22.5667 | -102.0917
22.5704 | -102.0753
22.5741 | -102.06
22.5778 | -102.0462
22.5815 | -102.032
22.5852 | -102.0179
22.5889 | -102.0039
22.5926 | -101.9937
22.5963 | -101.9834
22.6 | -101.9744
22.6037 | -101.9651
22.6074 | -101.9565
22.6111 | -101.9475
22.6148 | -101.9427
22.6185 | -101.9413
22.6222 | -101.9393
22.6259 | -101.9385
22.6296 | -101.9409
22.6333 | -101.9455
22.637 | -101.9504
22.6407 | -101.956
22.6444 | -101.9644
22.6481 | -101.9741
22.6519 | -101.9853
22.6556 | -101.9977
22.6593 | -102.0124
22.663 | -102.0291
22.6667 | -102.0476
22.6704 | -102.0689
22.6741 | -102.0888
22.6778 | -102.1089
22.6815 | -102.1306
22.6852 | -102.154
22.6889 | -102.1793
22.6926 | -102.2035
22.6963 | -102.2278
22.7 | -102.2532
22.7037 | -102.2776
22.7074 | -102.3027
22.7111 | -102.3285
22.7148 | -102.3543
22.7185 | -102.3813
22.7222 | -102.4093
22.7259 | -102.4386
22.7296 | -102.4657
22.7333 | -102.4933
22.737 | -102.5205
22.7407 | -102.5448
22.7444 | -102.5661
22.7481 | -102.5871
22.7519 | -102.6104
22.7556 | -102.6331
22.7593 | -102.6541
22.763 | -102.6752
22.7667 | -102.6965
22.7704 | -102.716
22.7741 | -102.7354
22.7778 | -102.752
22.7815 | -102.7669
22.7852 | -102.7801
22.7889 | -102.7922
22.7926 | -102.8029
22.7963 | -102.8116
22.8 | -102.8186
22.8037 | -102.8231
22.8074 | -102.8266
22.8111 | -102.829
22.8148 | -102.8326
22.8185 | -102.8345
22.8222 | -102.8327
22.8259 | -102.8315
22.8296 | -102.8303
22.8333 | -102.8291
22.837 | -102.8282
22.8407 | -102.8265
22.8444 | -102.8236
22.8481 | -102.8203
22.8519 | -102.8172
22.8556 | -102.811
22.8593 | -102.8057
22.863 | -102.8038
22.8667 | -102.7992
22.8704 | -102.7957
22.8741 | -102.7903
22.8778 | -102.7835
22.8815 | -102.7788
22.8852 | -102.7749
22.8889 | -102.7723
22.8926 | -102.7674
22.8963 | -102.7621
22.9 | -102.7567
22.9037 | -102.7527
22.9074 | -102.7487
22.9111 | -102.7452
22.9148 | -102.7415
22.9185 | -102.7377
22.9222 | -102.7358
22.9259 | -102.7338
22.9296 | -102.734
22.9333 | -102.7358
22.937 | -102.7349
22.9407 | -102.7323
22.9444 | -102.7326
22.9481 | -102.733
22.9519 | -102.7311
22.9556 | -102.7298
22.9593 | -102.7302
22.963 | -102.7292
22.9667 | -102.7273
22.9704 | -102.7254
22.9741 | -102.7239
22.9778 | -102.7221
22.9815 | -102.7195
22.9852 | -102.7173
22.9889 | -102.7157
22.9926 | -102.7152
22.9963 | -102.7138
23 | -102.7124
23.0037 | -102.7108
23.0074 | -102.7092
23.0111 | -102.7072
23.0148 | -102.7035
23.0185 | -102.7015
23.0222 | -102.6996
23.0259 | -102.6976
23.0296 | -102.6949
23.0333 | -102.6901
23.037 | -102.6869
23.0407 | -102.6835
23.0444 | -102.6804
23.0481 | -102.6778
23.0519 | -102.6746
23.0556 | -102.6729
23.0593 | -102.6713
23.063 | -102.6708
23.0667 | -102.6713
23.0704 | -102.6736
23.0741 | -102.6751
23.0778 | -102.6762
23.0815 | -102.6785
23.0852 | -102.6799
23.0889 | -102.6813
23.0926 | -102.6834
23.0963 | -102.6878
23.1 | -102.6928
23.1037 | -102.7003
23.1074 | -102.7087
23.1111 | -102.7168
23.1148 | -102.7242
23.1185 | -102.7338
23.1222 | -102.7454
23.1259 | -102.755
23.1296 | -102.7673
23.1333 | -102.7795
23.137 | -102.7916
23.1407 | -102.8039
23.1444 | -102.8162
23.1481 | -102.8303
23.1519 | -102.8439
23.1556 | -102.8554
23.1593 | -102.8685
23.163 | -102.8808
23.1667 | -102.893
23.1704 | -102.9048
23.1741 | -102.9158
23.1778 | -102.9261
23.1815 | -102.9358
23.1852 | -102.9433
23.1889 | -102.9525
23.1926 | -102.9603
23.1963 | -102.9643
23.2 | -102.9716
23.2037 | -102.9778
23.2074 | -102.9823
23.2111 | -102.9853
23.2148 | -102.9873
23.2185 | -102.9888
23.2222 | -102.9901
23.2259 | -102.9879
23.2296 | -102.9838
23.2333 | -102.9794
23.237 | -102.9737
23.2407 | -102.9653
23.2444 | -102.9582
23.2481 | -102.9493
23.2519 | -102.9413
23.2556 | -102.9295
23.2593 | -102.9164
23.263 | -102.9037
23.2667 | -102.8925
23.2704 | -102.8785
23.2741 | -102.8666
23.2778 | -102.8557
23.2815 | -102.8429
23.2852 | -102.8301
23.2889 | -102.8191
23.2926 | -102.8079
23.2963 | -102.7967
23.3 | -102.7818
23.3037 | -102.767
23.3074 | -102.7532
23.3111 | -102.7411
23.3148 | -102.7289
23.3185 | -102.7152
23.3222 | -102.7022
23.3259 | -102.6886
23.3296 | -102.675
23.3333 | -102.663
23.337 | -102.6519
23.3407 | -102.64
23.3444 | -102.6291
23.3481 | -102.617
23.3519 | -102.6071
23.3556 | -102.598
23.3593 | -102.5895
23.363 | -102.5818
23.3667 | -102.5734
23.3704 | -102.5682
23.3741 | -102.5611
23.3778 | -102.5534
23.3815 | -102.5485
23.3852 | -102.5436
23.3889 | -102.5389
23.3926 | -102.5339
23.3963 | -102.5291
23.4 | -102.5253
23.4037 | -102.5233
23.4074 | -102.5214
23.4111 | -102.5207
23.4148 | -102.5211
23.4185 | -102.5222
23.4222 | -102.5219
23.4259 | -102.521
23.4296 | -102.5201
23.4333 | -102.5196
23.437 | -102.5187
23.4407 | -102.5193
23.4444 | -102.5211
23.4481 | -102.5226
23.4519 | -102.5243
23.4556 | -102.5271
23.4593 | -102.5291
23.463 | -102.5294
23.4667 | -102.5304
23.4704 | -102.535
23.4741 | -102.5408
23.4778 | -102.5474
23.4815 | -102.5533
23.4852 | -102.561
23.4889 | -102.5674
23.4926 | -102.5738
23.4963 | -102.5799
23.5 | -102.5877
23.5037 | -102.596
23.5074 | -102.6064
23.5111 | -102.6157
23.5148 | -102.6259
23.5185 | -102.6355
23.5222 | -102.6431
23.5259 | -102.6514
23.5296 | -102.6605
23.5333 | -102.6689
23.537 | -102.6779
23.5407 | -102.687
23.5444 | -102.6966
23.5481 | -102.7069
23.5519 | -102.7168
23.5556 | -102.7262
23.5593 | -102.7374
23.563 | -102.7482
23.5667 | -102.758
23.5704 | -102.7669
23.5741 | -102.7754
23.5778 | -102.7845
23.5815 | -102.7935
23.5852 | -102.8022
23.5889 | -102.8116
23.5926 | -102.8202
23.5963 | -102.8274
23.6 | -102.8337
23.6037 | -102.8411
23.6074 | -102.8475
23.6111 | -102.8525
23.6148 | -102.8581
23.6185 | -102.863
23.6222 | -102.8672
23.6259 | -102.8712
23.6296 | -102.8757
23.6333 | -102.8804
23.637 | -102.8845
23.6407 | -102.8899
23.6444 | -102.8941
23.6481 | -102.8964
23.6519 | -102.8987
23.6556 | -102.9009
23.6593 | -102.9038
23.663 | -102.908
23.6667 | -102.9123
23.6704 | -102.9126
23.6741 | -102.9136
23.6778 | -102.9144
23.6815 | -102.9155
23.6852 | -102.9127
23.6889 | -102.9095
23.6926 | -102.907
23.6963 | -102.9032
23.7 | -102.9014
23.7037 | -102.8998
23.7074 | -102.8992
23.7111 | -102.899
23.7148 | -102.8976
23.7185 | -102.8969
23.7222 | -102.8985
23.7259 | -102.897
23.7296 | -102.897
23.7333 | -102.8965
23.737 | -102.8965
23.7407 | -102.898
23.7444 | -102.8996
23.7481 | -102.9019
23.7519 | -102.9049
23.7556 | -102.9078
23.7593 | -102.9113
23.763 | -102.9148
23.7667 | -102.916
23.7704 | -102.9213
23.7741 | -102.9276
23.7778 | -102.9306
23.7815 | -102.9344
23.7852 | -102.9398
23.7889 | -102.9442
23.7926 | -102.9475
23.7963 | -102.9551
23.8 | -102.9648
23.8037 | -102.9704
23.8074 | -102.9748
23.8111 | -102.9806
23.8148 | -102.9868
23.8185 | -102.9949
23.8222 | -103.0032
23.8259 | -103.0122
23.8296 | -103.0216
23.8333 | -103.0293
23.837 | -103.0379
23.8407 | -103.0474
23.8444 | -103.0557
23.8481 | -103.0637
23.8519 | -103.0727
23.8556 | -103.0809
23.8593 | -103.0889
23.863 | -103.0988
23.8667 | -103.1069
23.8704 | -103.1139
23.8741 | -103.1212
23.8778 | -103.1272
23.8815 | -103.1346
23.8852 | -103.1407
23.8889 | -103.1486
23.8926 | -103.1556
23.8963 | -103.1621
23.9 | -103.1687
23.9037 | -103.1752
23.9074 | -103.1817
23.9111 | -103.188
23.9148 | -103.1913
23.9185 | -103.1937
23.9222 | -103.1973
23.9259 | -103.2006
23.9296 | -103.2048
23.9333 | -103.2078
23.937 | -103.2094
23.9407 | -103.2107
23.9444 | -103.2124
23.9481 | -103.215
23.9519 | -103.2174
23.9556 | -103.2167
23.9593 | -103.2178
23.963 | -103.2207
23.9667 | -103.223
23.9704 | -103.2268
23.9741 | -103.2303
23.9778 | -103.2314
23.9815 | -103.232
23.9852 | -103.2321
23.9889 | -103.2339
23.9926 | -103.2348
23.9963 | -103.2356
24 | -103.2362
24.0037 | -103.2365
24.0074 | -103.2399
24.0111 | -103.241
24.0148 | -103.2394
24.0185 | -103.2387
24.0222 | -103.2422
24.0259 | -103.245
24.0296 | -103.2472
24.0333 | -103.2507
24.037 | -103.2507
24.0407 | -103.2506
24.0444 | -103.2499
24.0481 | -103.2482
24.0519 | -103.2477
24.0556 | -103.2464
24.0593 | -103.2433
24.063 | -103.2413
24.0667 | -103.2411
24.0704 | -103.2391
24.0741 | -103.2373
24.0778 | -103.2337
24.0815 | -103.2308
24.0852 | -103.2263
24.0889 | -103.2214
24.0926 | -103.2165
24.0963 | -103.2144
24.1 | -103.2122
24.1037 | -103.2103
24.1074 | -103.2096
24.1111 | -103.2049
24.1148 | -103.1982
24.1185 | -103.1931
24.1222 | -103.1861
24.1259 | -103.1761
24.1296 | -103.1673
24.1333 | -103.1596
24.137 | -103.1497
24.1407 | -103.1395
24.1444 | -103.1285
24.1481 | -103.1177
24.1519 | -103.1058
24.1556 | -103.0958
24.1593 | -103.0814
24.163 | -103.0672
24.1667 | -103.0523
24.1704 | -103.0391
24.1741 | -103.0264
24.1778 | -103.0135
24.1815 | -103.0009
24.1852 | -102.987
24.1889 | -102.9744
24.1926 | -102.9612
24.1963 | -102.9487
24.2 | -102.9361
24.2037 | -102.9218
24.2074 | -102.9104
24.2111 | -102.8988
24.2148 | -102.8866
24.2185 | -102.8753
24.2222 | -102.8607
24.2259 | -102.847
24.2296 | -102.8334
24.2333 | -102.8216
24.237 | -102.8115
24.2407 | -102.8018
24.2444 | -102.7912
24.2481 | -102.7808
24.2519 | -102.7721
24.2556 | -102.7649
24.2593 | -102.7587
24.263 | -102.749
24.2667 | -102.7246
24.2704 | -102.6777
24.2741 | -102.6151
24.2778 | -102.5559
24.2815 | -102.511
24.2852 | -102.473
24.2889 | -102.437
24.2926 | -102.4004
24.2963 | -102.3659
24.3 | -102.3287
24.3037 | -102.2883
24.3074 | -102.2469
24.3111 | -102.2042
24.3148 | -102.1588
24.3185 | -102.1106
24.3222 | -102.0617
24.3259 | -102.0121
24.3296 | -101.9585
24.3333 | -101.9034
24.337 | -101.8467
24.3407 | -101.7886
24.3444 | -101.7297
24.3481 | -101.6687
24.3519 | -101.608
24.3556 | -101.5458
24.3593 | -101.4829
24.363 | -101.42
24.3667 | -101.3582
24.3704 | -101.2962
24.3741 | -101.2347
24.3778 | -101.1746
24.3815 | -101.1146
24.3852 | -101.0563
24.3889 | -101.0022
24.3926 | -100.9492
24.3963 | -100.8988
24.4 | -100.8521
24.4037 | -100.8087
24.4074 | -100.7672
24.4111 | -100.7296
24.4148 | -100.6959
24.4185 | -100.665
24.4222 | -100.6379
24.4259 | -100.6144
24.4296 | -100.5971
24.4333 | -100.5841
24.437 | -100.5755
24.4407 | -100.5719
24.4444 | -100.5723
24.4481 | -100.5759
24.4519 | -100.5851
24.4556 | -100.6005
24.4593 | -100.621
24.463 | -100.6453
24.4667 | -100.673
24.4704 | -100.7041
24.4741 | -100.7396
24.4778 | -100.7793
24.4815 | -100.8209
24.4852 | -100.8653
24.4889 | -100.9154
24.4926 | -100.9686
24.4963 | -101.0236
24.5 | -101.0797
24.5037 | -101.1409
24.5074 | -101.2041
24.5111 | -101.2685
24.5148 | -101.3329
24.5185 | -101.3989
24.5222 | -101.4649
24.5259 | -101.5304
24.5296 | -101.5971
24.5333 | -101.6618
24.537 | -101.7254
24.5407 | -101.7889
24.5444 | -101.8528
24.5481 | -101.9157
24.5519 | -101.9739
24.5556 | -102.031
24.5593 | -102.0869
24.563 | -102.1389
24.5667 | -102.1855
24.5704 | -102.2273
24.5741 | -102.2578
24.5778 | -102.2111
24.5815 | -102.2105
24.5852 | -102.198
24.5889 | -102.1724
24.5926 | -102.1454
24.5963 | -102.1188
24.6 | -102.0946
24.6037 | -102.066
24.6074 | -102.0298
24.6111 | -101.9893
24.6148 | -101.9506
24.6185 | -101.92
24.6222 | -101.8935
24.6259 | -101.8525
24.6296 | -101.7901
24.6333 | -101.715
24.637 | -101.6342
24.6407 | -101.5487
24.6444 | -101.4572
24.6481 | -101.3633
24.6519 | -101.2678
24.6556 | -101.1701
24.6593 | -101.0707
24.663 | -100.973
24.6667 | -100.8754
24.6704 | -100.777
24.6741 | -100.6785
24.6778 | -100.5812
24.6815 | -100.4843
24.6852 | -100.388
24.6889 | -100.2941
24.6926 | -100.2028
24.6963 | -100.1136
24.7 | -100.0271
24.7037 | -99.9448
24.7074 | -99.8653
24.7111 | -99.7886
24.7148 | -99.7151
24.7185 | -99.6464
24.7222 | -99.5829
24.7259 | -99.5243
24.7296 | -99.4706
24.7333 | -99.4226
24.737 | -99.3786
24.7407 | -99.3403
24.7444 | -99.3069
24.7481 | -99.2786
24.7519 | -99.2564
24.7556 | -99.2403
24.7593 | -99.2304
24.763 | -99.2259
24.7667 | -99.2267
24.7704 | -99.2352
24.7741 | -99.25
24.7778 | -99.2692
24.7815 | -99.2941
24.7852 | -99.3253
24.7889 | -99.3611
24.7926 | -99.4031
24.7963 | -99.4516
24.8 | -99.505
24.8037 | -99.563
24.8074 | -99.626
24.8111 | -99.6949
24.8148 | -99.7696
24.8185 | -99.8498
24.8222 | -99.9328
24.8259 | -100.0191
24.8296 | -100.11
24.8333 | -100.2031
24.837 | -100.2975
24.8407 | -100.3962
24.8444 | -100.498
24.8481 | -100.6019
24.8519 | -100.7064
24.8556 | -100.8116
24.8593 | -100.9176
24.863 | -101.0251
24.8667 | -101.1329
24.8704 | -101.2401
24.8741 | -101.3469
24.8778 | -101.4525
24.8815 | -101.5583
24.8852 | -101.662
24.8889 | -101.7628
24.8926 | -101.8588
24.8963 | -101.952
24.9 | -102.0449
24.9037 | -102.1337
24.9074 | -102.219
24.9111 | -102.303
24.9148 | -102.3865
24.9185 | -102.4656
24.9222 | -102.5477
24.9259 | -102.6336
24.9296 | -102.7232
24.9333 | -102.8026
24.937 | -102.8809
24.9407 | -102.9592
24.9444 | -103.0246
24.9481 | -103.0749
24.9519 | -103.1091
24.9556 | -103.1367
24.9593 | -103.1635
24.963 | -103.1893
24.9667 | -103.2143
24.9704 | -103.2348
24.9741 | -103.2488
24.9778 | -103.2597
24.9815 | -103.2663
24.9852 | -103.2722
24.9889 | -103.2771
24.9926 | -103.2826
24.9963 | -103.29
25 | -103.2979
25.0037 | -103.3056
25.0074 | -103.3127
25.0111 | -103.321
25.0148 | -103.327
25.0185 | -103.3336
25.0222 | -103.3411
25.0259 | -103.3494
25.0296 | -103.3577
25.0333 | -103.3651
25.037 | -103.374
25.0407 | -103.3824
25.0444 | -103.3894
25.0481 | -103.3992
25.0519 | -103.4101
25.0556 | -103.4206
25.0593 | -103.4321
25.063 | -103.4427
25.0667 | -103.4539
25.0704 | -103.4658
25.0741 | -103.478
25.0778 | -103.4889
25.0815 | -103.4983
25.0852 | -103.5088
25.0889 | -103.5207
25.0926 | -103.5323
25.0963 | -103.5462
25.1 | -103.562
25.1037 | -103.5758
25.1074 | -103.5879
25.1111 | -103.5974
25.1148 | -103.6069
25.1185 | -103.6184
25.1222 | -103.6314
25.1259 | -103.6408
25.1296 | -103.6488
25.1333 | -103.6581
25.137 | -103.6677
25.1407 | -103.6757
25.1444 | -103.6849
25.1481 | -103.6943
25.1519 | -103.7031
25.1556 | -103.7096
25.1593 | -103.7172
25.163 | -103.7255
25.1667 | -103.7325
25.1704 | -103.739
25.1741 | -103.7461
25.1778 | -103.7543
25.1815 | -103.7606
25.1852 | -103.7669
25.1889 | -103.7752
25.1926 | -103.783
25.1963 | -103.7871
25.2 | -103.7917
25.2037 | -103.7968
25.2074 | -103.8023
25.2111 | -103.8096
25.2148 | -103.8162
25.2185 | -103.8206
25.2222 | -103.8244
25.2259 | -103.8294
25.2296 | -103.8347
25.2333 | -103.8404
25.237 | -103.8443
25.2407 | -103.8493
25.2444 | -103.8543
25.2481 | -103.8607
25.2519 | -103.8666
25.2556 | -103.8709
25.2593 | -103.8756
25.263 | -103.879
25.2667 | -103.8831
25.2704 | -103.8862
25.2741 | -103.8903
25.2778 | -103.8936
25.2815 | -103.8957
25.2852 | -103.8982
25.2889 | -103.9029
25.2926 | -103.9062
25.2963 | -103.9088
25.3 | -103.9107
25.3037 | -103.9136
25.3074 | -103.9153
25.3111 | -103.9168
25.3148 | -103.9182
25.3185 | -103.9188
25.3222 | -103.9222
25.3259 | -103.9252
25.3296 | -103.9257
25.3333 | -103.9272
25.337 | -103.9302
25.3407 | -103.9321
25.3444 | -103.9334
25.3481 | -103.9352
25.3519 | -103.9372
25.3556 | -103.9378
25.3593 | -103.939
25.363 | -103.9418
25.3667 | -103.9428
25.3704 | -103.9432
25.3741 | -103.9438
25.3778 | -103.9434
25.3815 | -103.9427
25.3852 | -103.9434
25.3889 | -103.9454
25.3926 | -103.9463
25.3963 | -103.9457
25.4 | -103.9472
25.4037 | -103.9464
25.4074 | -103.9442
25.4111 | -103.9431
25.4148 | -103.9431
25.4185 | -103.9431
25.4222 | -103.9433
25.4259 | -103.9437
25.4296 | -103.9434
25.4333 | -103.9445
25.437 | -103.9472
25.4407 | -103.9484
25.4444 | -103.9489
25.4481 | -103.95
25.4519 | -103.9471
25.4556 | -103.9427
25.4593 | -103.9404
25.463 | -103.941
25.4667 | -103.9404
25.4704 | -103.9405
25.4741 | -103.9415
25.4778 | -103.9403
25.4815 | -103.9407
25.4852 | -103.9409
25.4889 | -103.9408
25.4926 | -103.9402
25.4963 | -103.9402
25.5 | -103.9386
25.5037 | -103.9366
25.5074 | -103.9353
25.5111 | -103.9339
25.5148 | -103.9325
25.5185 | -103.9301
25.5222 | -103.9276
25.5259 | -103.9254
25.5296 | -103.9228
25.5333 | -103.9193
25.537 | -103.9175
25.5407 | -103.9171
25.5444 | -103.9152
25.5481 | -103.9121
25.5519 | -103.911
25.5556 | -103.9091
25.5593 | -103.9067
25.563 | -103.9043
25.5667 | -103.9012
25.5704 | -103.8963
25.5741 | -103.8937
25.5778 | -103.8924
25.5815 | -103.8876
25.5852 | -103.8823
25.5889 | -103.8779
25.5926 | -103.8726
25.5963 | -103.8688
25.6 | -103.8642
25.6037 | -103.8613
25.6074 | -103.8592
25.6111 | -103.8565
25.6148 | -103.8539
25.6185 | -103.8504
25.6222 | -103.8479
25.6259 | -103.8449
25.6296 | -103.8448
25.6333 | -103.8448
25.637 | -103.8451
25.6407 | -103.8453
25.6444 | -103.8442
25.6481 | -103.8436
25.6519 | -103.8414
25.6556 | -103.8376
25.6593 | -103.8364
25.663 | -103.8351
25.6667 | -103.8322
25.6704 | -103.8315
25.6741 | -103.8302
25.6778 | -103.83
25.6815 | -103.8313
25.6852 | -103.8314
25.6889 | -103.8324
25.6926 | -103.8344
25.6963 | -103.8347
25.7 | -103.8335
25.7037 | -103.8338
25.7074 | -103.8322
25.7111 | -103.8315
25.7148 | -103.8313
25.7185 | -103.8316
25.7222 | -103.8322
25.7259 | -103.8326
25.7296 | -103.8353
25.7333 | -103.8386
25.737 | -103.839
25.7407 | -103.8397
25.7444 | -103.8409
25.7481 | -103.8416
25.7519 | -103.8416
25.7556 | -103.8439
25.7593 | -103.8462
25.763 | -103.8444
25.7667 | -103.8444
25.7704 | -103.8463
25.7741 | -103.8466
25.7778 | -103.8456
25.7815 | -103.8449
25.7852 | -103.8452
25.7889 | -103.8447
25.7926 | -103.8447
25.7963 | -103.8446
25.8 | -103.844
25.8037 | -103.8436
25.8074 | -103.8439
25.8111 | -103.8439
25.8148 | -103.8468
25.8185 | -103.8494
25.8222 | -103.8518
25.8259 | -103.8522
25.8296 | -103.8526
25.8333 | -103.8536
25.837 | -103.8558
25.8407 | -103.8587
25.8444 | -103.8586
25.8481 | -103.8599
25.8519 | -103.8624
25.8556 | -103.8641
25.8593 | -103.8671
25.863 | -103.8716
25.8667 | -103.8763
25.8704 | -103.8783
25.8741 | -103.8825
25.8778 | -103.8893
25.8815 | -103.8953
25.8852 | -103.9012
25.8889 | -103.9074
25.8926 | -103.9144
25.8963 | -103.9221
25.9 | -103.9283
25.9037 | -103.9334
25.9074 | -103.9386
25.9111 | -103.9444
25.9148 | -103.9498
25.9185 | -103.9566
25.9222 | -103.9637
25.9259 | -103.9699
25.9296 | -103.9763
25.9333 | -103.9855
25.937 | -103.9928
25.9407 | -103.9987
25.9444 | -104.0069
25.9481 | -104.0136
25.9519 | -104.0183
25.9556 | -104.023
25.9593 | -104.0278
25.963 | -104.034
25.9667 | -104.0366
25.9704 | -104.0406
25.9741 | -104.0452
25.9778 | -104.0481
25.9815 | -104.0509
25.9852 | -104.0531
25.9889 | -104.055
25.9926 | -104.0557
25.9963 | -104.0567
26 | -104.0556
26.0037 | -104.056
26.0074 | -104.0589
26.0111 | -104.0609
26.0148 | -104.0626
26.0185 | -104.0631
26.0222 | -104.0646
26.0259 | -104.0625
26.0296 | -104.0601
26.0333 | -104.0594
26.037 | -104.0553
26.0407 | -104.0519
26.0444 | -104.0482
26.0481 | -104.0435
26.0519 | -104.0383
26.0556 | -104.0337
26.0593 | -104.0304
26.063 | -104.027
26.0667 | -104.0231
26.0704 | -104.0165
26.0741 | -104.0105
26.0778 | -104.0027
26.0815 | -103.9955
26.0852 | -103.991
26.0889 | -103.9854
26.0926 | -103.9775
26.0963 | -103.9684
26.1 | -103.9612
26.1037 | -103.9569
26.1074 | -103.9507
26.1111 | -103.9446
26.1148 | -103.9387
26.1185 | -103.9333
26.1222 | -103.9263
26.1259 | -103.9196
26.1296 | -103.9147
26.1333 | -103.9107
26.137 | -103.9036
26.1407 | -103.8974
26.1444 | -103.8924
26.1481 | -103.8882
26.1519 | -103.8833
26.1556 | -103.8785
26.1593 | -103.8742
26.163 | -103.8699
26.1667 | -103.8661
26.1704 | -103.862
26.1741 | -103.8589
26.1778 | -103.8581
26.1815 | -103.8615
26.1852 | -103.8605
26.1889 | -103.8607
26.1926 | -103.8609
26.1963 | -103.8631
26.2 | -103.8656
26.2037 | -103.8673
26.2074 | -103.8683
26.2111 | -103.87
26.2148 | -103.8746
26.2185 | -103.879
26.2222 | -103.8837
26.2259 | -103.89
26.2296 | -103.8952
26.2333 | -103.9009
26.237 | -103.9052
26.2407 | -103.9115
26.2444 | -103.9181
26.2481 | -103.9238
26.2519 | -103.9305
26.2556 | -103.9382
26.2593 | -103.9466
26.263 | -103.9562
26.2667 | -103.966
26.2704 | -103.9768
26.2741 | -103.9891
26.2778 | -104.0012
26.2815 | -104.0124
26.2852 | -104.0241
26.2889 | -104.0351
26.2926 | -104.0443
26.2963 | -104.0555
26.3 | -104.067
26.3037 | -104.0759
26.3074 | -104.0858
26.3111 | -104.0965
26.3148 | -104.1081
26.3185 | -104.1209
26.3222 | -104.1307
26.3259 | -104.1393
26.3296 | -104.1487
26.3333 | -104.1567
26.337 | -104.1661
26.3407 | -104.1735
26.3444 | -104.1805
26.3481 | -104.1896
26.3519 | -104.1987
26.3556 | -104.2067
26.3593 | -104.214
26.363 | -104.2209
26.3667 | -104.2269
26.3704 | -104.2328
26.3741 | -104.2378
26.3778 | -104.2419
26.3815 | -104.2458
26.3852 | -104.2481
26.3889 | -104.2504
26.3926 | -104.2543
26.3963 | -104.2588
26.4 | -104.2627
26.4037 | -104.2655
26.4074 | -104.2691
26.4111 | -104.2721
26.4148 | -104.2742
26.4185 | -104.2746
26.4222 | -104.2755
26.4259 | -104.2752
26.4296 | -104.2746
26.4333 | -104.2743
26.437 | -104.2747
26.4407 | -104.2749
26.4444 | -104.2759
26.4481 | -104.277
26.4519 | -104.2759
26.4556 | -104.2711
26.4593 | -104.2681
26.463 | -104.268
26.4667 | -104.2683
26.4704 | -104.2657
26.4741 | -104.2629
26.4778 | -104.2618
26.4815 | -104.2604
26.4852 | -104.2581
26.4889 | -104.2552
26.4926 | -104.2525
26.4963 | -104.2502
26.5 | -104.2473
26.5037 | -104.2439
26.5074 | -104.2434
26.5111 | -104.2407
26.5148 | -104.2369
26.5185 | -104.2332
26.5222 | -104.2294
26.5259 | -104.2276
26.5296 | -104.2238
26.5333 | -104.2205
26.537 | -104.219
26.5407 | -104.2173
26.5444 | -104.2146
26.5481 | -104.2092
26.5519 | -104.2057
26.5556 | -104.2008
26.5593 | -104.1965
26.563 | -104.1933
26.5667 | -104.1907
26.5704 | -104.1876
26.5741 | -104.1855
26.5778 | -104.1828
26.5815 | -104.1801
26.5852 | -104.1785
26.5889 | -104.1773
26.5926 | -104.1748
26.5963 | -104.1712
26.6 | -104.1687
26.6037 | -104.1673
26.6074 | -104.1653
26.6111 | -104.1649
26.6148 | -104.1647
26.6185 | -104.1639
26.6222 | -104.1633
26.6259 | -104.1627
26.6296 | -104.1624
26.6333 | -104.1606
26.637 | -104.158
26.6407 | -104.157
26.6444 | -104.1562
26.6481 | -104.1561
26.6519 | -104.1574
26.6556 | -104.1606
26.6593 | -104.163
26.663 | -104.1657
26.6667 | -104.1703
26.6704 | -104.1735
26.6741 | -104.1763
26.6778 | -104.1789
26.6815 | -104.1809
26.6852 | -104.183
26.6889 | -104.1857
26.6926 | -104.1893
26.6963 | -104.1952
26.7 | -104.2006
26.7037 | -104.2048
26.7074 | -104.2084
26.7111 | -104.213
26.7148 | -104.2172
26.7185 | -104.2211
26.7222 | -104.2246
26.7259 | -104.2274
26.7296 | -104.2307
26.7333 | -104.2347
26.737 | -104.2378
26.7407 | -104.2406
26.7444 | -104.2426
26.7481 | -104.2454
26.7519 | -104.2476
26.7556 | -104.2497
26.7593 | -104.2522
26.763 | -104.255
26.7667 | -104.2578
26.7704 | -104.2579
26.7741 | -104.2588
26.7778 | -104.2616
26.7815 | -104.2631
26.7852 | -104.2652
26.7889 | -104.2675
26.7926 | -104.268
26.7963 | -104.2694
26.8 | -104.2695
26.8037 | -104.2696
26.8074 | -104.271
26.8111 | -104.2726
26.8148 | -104.2742
26.8185 | -104.2752
26.8222 | -104.2767
26.8259 | -104.2767
26.8296 | -104.2766
26.8333 | -104.2787
26.837 | -104.2796
26.8407 | -104.28
26.8444 | -104.2806
26.8481 | -104.2816
26.8519 | -104.2833
26.8556 | -104.2844
26.8593 | -104.2859
26.863 | -104.2873
26.8667 | -104.2885
26.8704 | -104.289
26.8741 | -104.2896
26.8778 | -104.2898
26.8815 | -104.2878
26.8852 | -104.2868
26.8889 | -104.289
26.8926 | -104.2914
26.8963 | -104.2936
26.9 | -104.2962
26.9037 | -104.2968
26.9074 | -104.2973
26.9111 | -104.2987
26.9148 | -104.2983
26.9185 | -104.2987
26.9222 | -104.3003
26.9259 | -104.3024
26.9296 | -104.3032
26.9333 | -104.3044
26.937 | -104.3046
26.9407 | -104.3048
26.9444 | -104.3061
26.9481 | -104.3064
26.9519 | -104.3065
26.9556 | -104.3082
26.9593 | -104.3107
26.963 | -104.3129
26.9667 | -104.3131
26.9704 | -104.3147
26.9741 | -104.3178
26.9778 | -104.3197
26.9815 | -104.3209
26.9852 | -104.3244
26.9889 | -104.3269
26.9926 | -104.3281
26.9963 | -104.3288
27 | -104.3299
27.0037 | -104.3306
27.0074 | -104.331
27.0111 | -104.3341
27.0148 | -104.3363
27.0185 | -104.3376
27.0222 | -104.3382
27.0259 | -104.338
27.0296 | -104.3378
27.0333 | -104.336
27.037 | -104.3345
27.0407 | -104.3332
27.0444 | -104.3324
27.0481 | -104.3324
27.0519 | -104.3328
27.0556 | -104.3332
27.0593 | -104.3333
27.063 | -104.3335
27.0667 | -104.333
27.0704 | -104.3324
27.0741 | -104.3326
27.0778 | -104.333
27.0815 | -104.3324
27.0852 | -104.3319
27.0889 | -104.3308
27.0926 | -104.3304
27.0963 | -104.3286
27.1 | -104.3279
27.1037 | -104.3285
27.1074 | -104.3296
27.1111 | -104.3295
27.1148 | -104.3299
27.1185 | -104.3315
27.1222 | -104.3321
27.1259 | -104.3328
27.1296 | -104.3313
27.1333 | -104.3301
27.137 | -104.3307
27.1407 | -104.331
27.1444 | -104.3313
27.1481 | -104.3318
27.1519 | -104.3319
27.1556 | -104.3323
27.1593 | -104.3315
27.163 | -104.3314
27.1667 | -104.3301
27.1704 | -104.3285
27.1741 | -104.3284
27.1778 | -104.329
27.1815 | -104.3295
27.1852 | -104.3307
27.1889 | -104.3301
27.1926 | -104.3283
27.1963 | -104.3271
27.2 | -104.325
27.2037 | -104.3223
27.2074 | -104.3208
27.2111 | -104.3187
27.2148 | -104.3137
27.2185 | -104.3109
27.2222 | -104.3067
27.2259 | -104.3016
27.2296 | -104.2969
27.2333 | -104.2909
27.237 | -104.2864
27.2407 | -104.2818
27.2444 | -104.2793
27.2481 | -104.2747
27.2519 | -104.271
27.2556 | -104.2686
27.2593 | -104.2647
27.263 | -104.2614
27.2667 | -104.2601
27.2704 | -104.2563
27.2741 | -104.2518
27.2778 | -104.2476
27.2815 | -104.2428
27.2852 | -104.2377
27.2889 | -104.2338
27.2926 | -104.2308
27.2963 | -104.2272
27.3 | -104.2227
27.3037 | -104.217
27.3074 | -104.2113
27.3111 | -104.2074
27.3148 | -104.2034
27.3185 | -104.1992
27.3222 | -104.1967
27.3259 | -104.193
27.3296 | -104.1874
27.3333 | -104.1826
27.337 | -104.1775
27.3407 | -104.1726
27.3444 | -104.1683
27.3481 | -104.1635
27.3519 | -104.1594
27.3556 | -104.1547
27.3593 | -104.1488
27.363 | -104.1436
27.3667 | -104.1391
27.3704 | -104.1327
27.3741 | -104.1259
27.3778 | -104.1214
27.3815 | -104.1178
27.3852 | -104.1135
27.3889 | -104.109
27.3926 | -104.104
27.3963 | -104.0985
27.4 | -104.0924
27.4037 | -104.0865
27.4074 | -104.0815
27.4111 | -104.0764
27.4148 | -104.0716
27.4185 | -104.0684
27.4222 | -104.0638
27.4259 | -104.0579
27.4296 | -104.0514
27.4333 | -104.0445
27.437 | -104.0395
27.4407 | -104.0349
27.4444 | -104.031
27.4481 | -104.0275
27.4519 | -104.0233
27.4556 | -104.0186
27.4593 | -104.0142
27.463 | -104.0094
27.4667 | -104.0067
27.4704 | -104.0046
27.4741 | -104.0026
27.4778 | -103.9991
27.4815 | -103.9975
27.4852 | -103.9972
27.4889 | -103.9965
27.4926 | -103.9956
27.4963 | -103.9955
27.5 | -103.9951
27.5037 | -103.9953
27.5074 | -103.9962
27.5111 | -103.9985
27.5148 | -104.0027
27.5185 | -104.0067
27.5222 | -104.0086
27.5259 | -104.0109
27.5296 | -104.0137
27.5333 | -104.0172
27.537 | -104.0222
27.5407 | -104.0265
27.5444 | -104.0301
27.5481 | -104.0341
27.5519 | -104.0398
27.5556 | -104.046
27.5593 | -104.0528
27.563 | -104.06
27.5667 | -104.0674
27.5704 | -104.0747
27.5741 | -104.0829
27.5778 | -104.0907
27.5815 | -104.098
27.5852 | -104.1064
27.5889 | -104.1135
27.5926 | -104.1198
27.5963 | -104.1278
27.6 | -104.1389
27.6037 | -104.1475
27.6074 | -104.1553
27.6111 | -104.1642
27.6148 | -104.172
27.6185 | -104.1785
27.6222 | -104.1853
27.6259 | -104.1919
27.6296 | -104.1969
27.6333 | -104.2024
27.637 | -104.2083
27.6407 | -104.2136
27.6444 | -104.2171
27.6481 | -104.221
27.6519 | -104.2237
27.6556 | -104.226
27.6593 | -104.2284
27.663 | -104.2292
27.6667 | -104.2299
27.6704 | -104.231
27.6741 | -104.2329
27.6778 | -104.2336
27.6815 | -104.2336
27.6852 | -104.233
27.6889 | -104.2335
27.6926 | -104.2348
27.6963 | -104.2347
27.7 | -104.2315
27.7037 | -104.2272
27.7074 | -104.223
27.7111 | -104.2177
27.7148 | -104.2127
27.7185 | -104.2079
27.7222 | -104.2024
27.7259 | -104.1979
27.7296 | -104.1922
27.7333 | -104.1841
27.737 | -104.1765
27.7407 | -104.1686
27.7444 | -104.1609
27.7481 | -104.1541
27.7519 | -104.147
27.7556 | -104.1389
27.7593 | -104.1302
27.763 | -104.1221
27.7667 | -104.1137
27.7704 | -104.1042
27.7741 | -104.0951
27.7778 | -104.0871
27.7815 | -104.0802
27.7852 | -104.0705
27.7889 | -104.0612
27.7926 | -104.055
27.7963 | -104.0493
27.8 | -104.0433
27.8037 | -104.0371
27.8074 | -104.0303
27.8111 | -104.0218
27.8148 | -104.0146
27.8185 | -104.0082
27.8222 | -104.0021
27.8259 | -103.9968
27.8296 | -103.9905
27.8333 | -103.9851
27.837 | -103.9813
27.8407 | -103.9776
27.8444 | -103.9741
27.8481 | -103.9692
27.8519 | -103.9647
27.8556 | -103.9621
27.8593 | -103.9602
27.863 | -103.9585
27.8667 | -103.956
27.8704 | -103.9557
27.8741 | -103.9531
27.8778 | -103.9507
27.8815 | -103.9496
27.8852 | -103.9503
27.8889 | -103.9515
27.8926 | -103.9538
27.8963 | -103.9562
27.9 | -103.9593
27.9037 | -103.9626
27.9074 | -103.9644
27.9111 | -103.9651
27.9148 | -103.9675
27.9185 | -103.9717
27.9222 | -103.9741
27.9259 | -103.9728
27.9296 | -103.972
27.9333 | -103.9702
27.937 | -103.9711
27.9407 | -103.9719
27.9444 | -103.9716
27.9481 | -103.9725
27.9519 | -103.9713
27.9556 | -103.9703
27.9593 | -103.9678
27.963 | -103.9641
27.9667 | -103.9601
27.9704 | -103.9571
27.9741 | -103.954
27.9778 | -103.9506
27.9815 | -103.9467
27.9852 | -103.944
27.9889 | -103.9403
27.9926 | -103.9327
27.9963 | -103.9244
28 | -103.9173
28.0037 | -103.9089