## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 239 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 253
Hood | 02-92B/P | 179
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1001.9 | hPa |  | PRC PRO 01
Humidity | 61.6 | %/Hr
Internal temperature of calibration bench | 24.23 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1009 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.5 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.98 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.99 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3098 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 30/05/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 05/06/2018 31/05/2018 05/06/2018
Technical Manager | Anthony HUE | Approval |  | 13/06/2018
CEO | Robert GLEMOT | Approval |  | 13/06/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4306 | 57.4554
0.012209 | -36.0018 | 38.2278
0.018313 | -34.9386 | 28.0815
0.024417 | -34.4797 | 21.8444
0.030521 | -34.3728 | 18.0883
0.036626 | -34.2312 | 15.0315
0.04273 | -34.1375 | 12.6571
0.048834 | -34.0928 | 11.1381
0.054938 | -34.028 | 9.9993
0.061043 | -33.9667 | 8.9111
0.067147 | -33.9574 | 8.0908
0.073251 | -33.9607 | 7.3304
0.079355 | -33.9819 | 6.6657
0.08546 | -33.9975 | 6.0647
0.091564 | -33.9931 | 5.6181
0.097668 | -33.9752 | 5.367
0.10377 | -33.9529 | 5.119
0.10988 | -33.9356 | 4.8437
0.11598 | -33.9269 | 4.5573
0.12209 | -33.9302 | 4.2833
0.12819 | -33.9422 | 4.0087
0.13429 | -33.9548 | 3.7248
0.1404 | -33.9585 | 3.4445
0.1465 | -33.9609 | 3.1863
0.15261 | -33.9584 | 2.9646
0.15871 | -33.9521 | 2.8009
0.16482 | -33.9445 | 2.6984
0.17092 | -33.9433 | 2.6632
0.17702 | -33.9422 | 2.5933
0.18313 | -33.9423 | 2.5161
0.18923 | -33.9426 | 2.4053
0.19534 | -33.9417 | 2.2742
0.20144 | -33.9403 | 2.126
0.20754 | -33.9369 | 2.002
0.21365 | -33.9357 | 1.8851
0.21975 | -33.935 | 1.7861
0.22586 | -33.936 | 1.6991
0.23196 | -33.9386 | 1.6284
0.23807 | -33.943 | 1.5676
0.24417 | -33.9474 | 1.5087
0.25027 | -33.9518 | 1.4457
0.25638 | -33.9546 | 1.3829
0.26248 | -33.9567 | 1.3146
0.26859 | -33.9579 | 1.2467
0.27469 | -33.9588 | 1.1736
0.2808 | -33.9605 | 1.1007
0.2869 | -33.962 | 1.0297
0.293 | -33.9635 | 0.96399
0.29911 | -33.9651 | 0.90669
0.30521 | -33.9667 | 0.85114
0.31132 | -33.9679 | 0.81674
0.31742 | -33.9696 | 0.77752
0.32353 | -33.9703 | 0.74588
0.32963 | -33.9713 | 0.72011
0.33573 | -33.9725 | 0.70261
0.34184 | -33.9738 | 0.6824
0.34794 | -33.9757 | 0.65593
0.35405 | -33.9777 | 0.63018
0.36015 | -33.98 | 0.60169
0.36626 | -33.9818 | 0.56987
0.37236 | -33.9838 | 0.52797
0.37846 | -33.9859 | 0.47658
0.38457 | -33.9871 | 0.41009
0.39067 | -33.9883 | 0.35637
0.39678 | -33.9893 | 0.29686
0.40288 | -33.9901 | 0.23577
0.40899 | -33.9904 | 0.17743
0.41509 | -33.9903 | 0.12116
0.42119 | -33.9906 | 0.078565
0.4273 | -33.9909 | 0.031068
0.4334 | -33.9911 | -0.015238
0.43951 | -33.9915 | -0.05721
0.44561 | -33.9923 | -0.094044
0.45172 | -33.9936 | -0.11971
0.45782 | -33.9943 | -0.15205
0.46392 | -33.9954 | -0.18338
0.47003 | -33.9963 | -0.21062
0.47613 | -33.997 | -0.23983
0.48224 | -33.9974 | -0.27065
0.48834 | -33.9964 | -0.30327
0.49445 | -33.9961 | -0.33604
0.50055 | -33.9956 | -0.36508
0.50665 | -33.9949 | -0.39703
0.51276 | -33.9939 | -0.42934
0.51886 | -33.9928 | -0.46191
0.52497 | -33.9915 | -0.48534
0.53107 | -33.9908 | -0.51379
0.53717 | -33.99 | -0.54121
0.54328 | -33.9891 | -0.56296
0.54938 | -33.9888 | -0.58442
0.55549 | -33.9887 | -0.6038
0.56159 | -33.9896 | -0.61415
0.5677 | -33.9896 | -0.62712
0.5738 | -33.9906 | -0.6426
0.5799 | -33.9913 | -0.65648
0.58601 | -33.9923 | -0.66517
0.59211 | -33.9941 | -0.67091
0.59822 | -33.9954 | -0.67992
0.60432 | -33.997 | -0.69107
0.61043 | -33.9983 | -0.69883
0.61653 | -33.9996 | -0.70947
0.62263 | -34.0014 | -0.71939
0.62874 | -34.0024 | -0.73179
0.63484 | -34.0035 | -0.74172
0.64095 | -34.0044 | -0.75291
0.64705 | -34.0051 | -0.76429
0.65316 | -34.0057 | -0.7785
0.65926 | -34.0062 | -0.79638
0.66536 | -34.0056 | -0.81604
0.67147 | -34.0056 | -0.83124
0.67757 | -34.0057 | -0.84757
0.68368 | -34.0054 | -0.86851
0.68978 | -34.0051 | -0.88736
0.69589 | -34.0046 | -0.9079
0.70199 | -34.004 | -0.93167
0.70809 | -34.0035 | -0.95228
0.7142 | -34.0033 | -0.97378
0.7203 | -34.0025 | -0.99375
0.72641 | -34.0022 | -1.0173
0.73251 | -34.0019 | -1.0381
0.73862 | -34.0023 | -1.0594
0.74472 | -34.0023 | -1.0798
0.75082 | -34.0023 | -1.1019
0.75693 | -34.0029 | -1.1233
0.76303 | -34.0037 | -1.145
0.76914 | -34.0058 | -1.1649
0.77524 | -34.0069 | -1.1842
0.78135 | -34.0084 | -1.2035
0.78745 | -34.01 | -1.2231
0.79355 | -34.0115 | -1.2421
0.79966 | -34.0135 | -1.2609
0.80576 | -34.0159 | -1.28
0.81187 | -34.0178 | -1.3013
0.81797 | -34.0198 | -1.3219
0.82408 | -34.0218 | -1.342
0.83018 | -34.0236 | -1.362
0.83628 | -34.0257 | -1.3801
0.84239 | -34.0271 | -1.4007
0.84849 | -34.0287 | -1.4229
0.8546 | -34.0304 | -1.4425
0.8607 | -34.0318 | -1.4631
0.86681 | -34.0331 | -1.4821
0.87291 | -34.0342 | -1.5013
0.87901 | -34.0347 | -1.5193
0.88512 | -34.0356 | -1.5375
0.89122 | -34.0364 | -1.5559
0.89733 | -34.0369 | -1.5712
0.90343 | -34.0374 | -1.5891
0.90953 | -34.0375 | -1.6038
0.91564 | -34.0375 | -1.6197
0.92174 | -34.0377 | -1.635
0.92785 | -34.0378 | -1.6498
0.93395 | -34.0378 | -1.6636
0.94006 | -34.038 | -1.6766
0.94616 | -34.0376 | -1.6885
0.95226 | -34.0377 | -1.7012
0.95837 | -34.0378 | -1.7153
0.96447 | -34.038 | -1.7284
0.97058 | -34.0378 | -1.7418
0.97668 | -34.0375 | -1.7551
0.98279 | -34.0373 | -1.7677
0.98889 | -34.0371 | -1.7783
0.99499 | -34.037 | -1.7919
1.0011 | -34.0369 | -1.8024
1.0072 | -34.0365 | -1.8137
1.0133 | -34.0362 | -1.8242
1.0194 | -34.0358 | -1.8365
1.0255 | -34.0356 | -1.8498
1.0316 | -34.035 | -1.8616
1.0377 | -34.0347 | -1.8752
1.0438 | -34.0342 | -1.8876
1.0499 | -34.0334 | -1.9027
1.056 | -34.0328 | -1.9165
1.0621 | -34.0322 | -1.9302
1.0682 | -34.0318 | -1.9425
1.0743 | -34.0312 | -1.9574
1.0805 | -34.0308 | -1.9715
1.0866 | -34.0299 | -1.9862
1.0927 | -34.029 | -2.0004
1.0988 | -34.0284 | -2.015
1.1049 | -34.0276 | -2.0286
1.111 | -34.0269 | -2.0434
1.1171 | -34.0262 | -2.0575
1.1232 | -34.0257 | -2.0723
1.1293 | -34.0254 | -2.086
1.1354 | -34.0251 | -2.1008
1.1415 | -34.0247 | -2.1161
1.1476 | -34.0241 | -2.1314
1.1537 | -34.0237 | -2.1461
1.1598 | -34.0238 | -2.1605
1.1659 | -34.0235 | -2.1732
1.172 | -34.0233 | -2.1867
1.1781 | -34.0231 | -2.2004
1.1842 | -34.0229 | -2.2145
1.1903 | -34.0233 | -2.2252
1.1964 | -34.0232 | -2.2384
1.2025 | -34.0233 | -2.252
1.2086 | -34.0236 | -2.264
1.2147 | -34.0237 | -2.2764
1.2209 | -34.024 | -2.2871
1.227 | -34.0248 | -2.2976
1.2331 | -34.0253 | -2.3095
1.2392 | -34.0256 | -2.321
1.2453 | -34.0261 | -2.3317
1.2514 | -34.0266 | -2.3413
1.2575 | -34.0273 | -2.352
1.2636 | -34.0282 | -2.3607
1.2697 | -34.0289 | -2.3697
1.2758 | -34.0297 | -2.3807
1.2819 | -34.0304 | -2.3904
1.288 | -34.0311 | -2.401
1.2941 | -34.0321 | -2.4094
1.3002 | -34.033 | -2.4176
1.3063 | -34.0336 | -2.4266
1.3124 | -34.0344 | -2.4354
1.3185 | -34.0352 | -2.4444
1.3246 | -34.0359 | -2.4534
1.3307 | -34.0369 | -2.462
1.3368 | -34.0378 | -2.4714
1.3429 | -34.0385 | -2.4807
1.349 | -34.0392 | -2.4901
1.3551 | -34.04 | -2.4998
1.3613 | -34.0406 | -2.5083
1.3674 | -34.0414 | -2.5192
1.3735 | -34.0419 | -2.5294
1.3796 | -34.0425 | -2.5387
1.3857 | -34.0429 | -2.5484
1.3918 | -34.0436 | -2.559
1.3979 | -34.0442 | -2.5687
1.404 | -34.0448 | -2.5806
1.4101 | -34.0453 | -2.5916
1.4162 | -34.0458 | -2.6027
1.4223 | -34.0464 | -2.6147
1.4284 | -34.047 | -2.6253
1.4345 | -34.0473 | -2.6375
1.4406 | -34.0477 | -2.6515
1.4467 | -34.0483 | -2.6637
1.4528 | -34.0488 | -2.6756
1.4589 | -34.0494 | -2.6881
1.465 | -34.0497 | -2.6993
1.4711 | -34.0503 | -2.7112
1.4772 | -34.0509 | -2.7249
1.4833 | -34.0515 | -2.7381
1.4894 | -34.0521 | -2.7505
1.4955 | -34.0527 | -2.7625
1.5016 | -34.0534 | -2.7752
1.5078 | -34.0538 | -2.789
1.5139 | -34.0543 | -2.8023
1.52 | -34.0548 | -2.815
1.5261 | -34.0552 | -2.8277
1.5322 | -34.0557 | -2.8416
1.5383 | -34.0561 | -2.8544
1.5444 | -34.0566 | -2.8665
1.5505 | -34.0572 | -2.88
1.5566 | -34.0578 | -2.8926
1.5627 | -34.0583 | -2.9062
1.5688 | -34.059 | -2.9198
1.5749 | -34.0594 | -2.932
1.581 | -34.0598 | -2.9446
1.5871 | -34.0602 | -2.9582
1.5932 | -34.0606 | -2.971
1.5993 | -34.0611 | -2.9849
1.6054 | -34.0614 | -2.9979
1.6115 | -34.0619 | -3.0119
1.6176 | -34.0622 | -3.0241
1.6237 | -34.0625 | -3.0374
1.6298 | -34.0627 | -3.0495
1.6359 | -34.0632 | -3.0618
1.642 | -34.0636 | -3.0733
1.6482 | -34.0638 | -3.0863
1.6543 | -34.0642 | -3.0997
1.6604 | -34.0645 | -3.1116
1.6665 | -34.0647 | -3.1238
1.6726 | -34.0649 | -3.1362
1.6787 | -34.0653 | -3.1484
1.6848 | -34.0653 | -3.1613
1.6909 | -34.0656 | -3.1739
1.697 | -34.0658 | -3.1859
1.7031 | -34.0659 | -3.1983
1.7092 | -34.0659 | -3.2103
1.7153 | -34.0661 | -3.2224
1.7214 | -34.066 | -3.2349
1.7275 | -34.066 | -3.2475
1.7336 | -34.066 | -3.2607
1.7397 | -34.0662 | -3.2732
1.7458 | -34.0663 | -3.2854
1.7519 | -34.0664 | -3.2979
1.758 | -34.0665 | -3.3104
1.7641 | -34.0667 | -3.3223
1.7702 | -34.0668 | -3.3355
1.7763 | -34.0668 | -3.3474
1.7824 | -34.0668 | -3.3603
1.7885 | -34.0669 | -3.3733
1.7947 | -34.0671 | -3.3857
1.8008 | -34.0673 | -3.3973
1.8069 | -34.0674 | -3.4092
1.813 | -34.0676 | -3.4219
1.8191 | -34.0677 | -3.4351
1.8252 | -34.0679 | -3.4483
1.8313 | -34.0682 | -3.461
1.8374 | -34.0685 | -3.4733
1.8435 | -34.0687 | -3.4856
1.8496 | -34.0688 | -3.4991
1.8557 | -34.0691 | -3.5123
1.8618 | -34.0694 | -3.5248
1.8679 | -34.0696 | -3.5373
1.874 | -34.0699 | -3.5497
1.8801 | -34.0702 | -3.5626
1.8862 | -34.0706 | -3.5749
1.8923 | -34.0708 | -3.5874
1.8984 | -34.071 | -3.5995
1.9045 | -34.0712 | -3.6115
1.9106 | -34.0714 | -3.6238
1.9167 | -34.0716 | -3.6362
1.9228 | -34.0718 | -3.6483
1.9289 | -34.072 | -3.6604
1.9351 | -34.0722 | -3.6714
1.9412 | -34.0725 | -3.683
1.9473 | -34.0729 | -3.6954
1.9534 | -34.0732 | -3.7071
1.9595 | -34.0734 | -3.719
1.9656 | -34.0738 | -3.7312
1.9717 | -34.0741 | -3.7431
1.9778 | -34.0745 | -3.7547
1.9839 | -34.0748 | -3.7661
1.99 | -34.0751 | -3.7772
1.9961 | -34.0753 | -3.7893
2.0022 | -34.0755 | -3.8005
2.0083 | -34.0756 | -3.8119
2.0144 | -34.0759 | -3.8231
2.0205 | -34.0763 | -3.8349
2.0266 | -34.0767 | -3.8462
2.0327 | -34.077 | -3.8578
2.0388 | -34.0774 | -3.8693
2.0449 | -34.0777 | -3.8807
2.051 | -34.078 | -3.8921
2.0571 | -34.0784 | -3.9033
2.0632 | -34.0786 | -3.9148
2.0693 | -34.0788 | -3.927
2.0754 | -34.0791 | -3.9381
2.0816 | -34.0794 | -3.9503
2.0877 | -34.0796 | -3.9616
2.0938 | -34.08 | -3.9737
2.0999 | -34.0802 | -3.9852
2.106 | -34.0804 | -3.9969
2.1121 | -34.0806 | -4.0088
2.1182 | -34.0809 | -4.0193
2.1243 | -34.0812 | -4.0299
2.1304 | -34.0814 | -4.0408
2.1365 | -34.0817 | -4.0526
2.1426 | -34.0819 | -4.0642
2.1487 | -34.0822 | -4.0753
2.1548 | -34.0827 | -4.0867
2.1609 | -34.0831 | -4.0978
2.167 | -34.0836 | -4.1088
2.1731 | -34.0839 | -4.1207
2.1792 | -34.0841 | -4.1313
2.1853 | -34.0845 | -4.1431
2.1914 | -34.0848 | -4.1543
2.1975 | -34.0851 | -4.1655
2.2036 | -34.0854 | -4.1774
2.2097 | -34.0857 | -4.1887
2.2158 | -34.0861 | -4.1989
2.222 | -34.0864 | -4.2091
2.2281 | -34.0866 | -4.2203
2.2342 | -34.087 | -4.2313
2.2403 | -34.0872 | -4.2416
2.2464 | -34.0875 | -4.2521
2.2525 | -34.0878 | -4.2634
2.2586 | -34.0881 | -4.2736
2.2647 | -34.0884 | -4.2837
2.2708 | -34.0887 | -4.2946
2.2769 | -34.0891 | -4.305
2.283 | -34.0893 | -4.3156
2.2891 | -34.0896 | -4.3261
2.2952 | -34.0898 | -4.3364
2.3013 | -34.0901 | -4.3468
2.3074 | -34.0904 | -4.3569
2.3135 | -34.0906 | -4.368
2.3196 | -34.0908 | -4.3778
2.3257 | -34.091 | -4.3882
2.3318 | -34.0913 | -4.399
2.3379 | -34.0915 | -4.4094
2.344 | -34.0917 | -4.4207
2.3501 | -34.0919 | -4.431
2.3562 | -34.0921 | -4.441
2.3623 | -34.0922 | -4.4509
2.3685 | -34.0924 | -4.4609
2.3746 | -34.0925 | -4.4717
2.3807 | -34.0928 | -4.4819
2.3868 | -34.093 | -4.4921
2.3929 | -34.0932 | -4.5024
2.399 | -34.0933 | -4.5126
2.4051 | -34.0936 | -4.5228
2.4112 | -34.0938 | -4.5331
2.4173 | -34.0939 | -4.5434
2.4234 | -34.094 | -4.5538
2.4295 | -34.0942 | -4.5644
2.4356 | -34.0944 | -4.5751
2.4417 | -34.0945 | -4.5861
2.4478 | -34.0947 | -4.5967
2.4539 | -34.0949 | -4.6066
2.46 | -34.0949 | -4.6178
2.4661 | -34.0951 | -4.6285
2.4722 | -34.0952 | -4.6384
2.4783 | -34.0953 | -4.6486
2.4844 | -34.0954 | -4.6593
2.4905 | -34.0955 | -4.6701
2.4966 | -34.0957 | -4.6807
2.5027 | -34.0958 | -4.6908
2.5089 | -34.0959 | -4.7014
2.515 | -34.096 | -4.7119
2.5211 | -34.0961 | -4.7224
2.5272 | -34.0961 | -4.7334
2.5333 | -34.0963 | -4.7443
2.5394 | -34.0964 | -4.7547
2.5455 | -34.0965 | -4.766
2.5516 | -34.0967 | -4.7761
2.5577 | -34.0969 | -4.7876
2.5638 | -34.097 | -4.7993
2.5699 | -34.097 | -4.8107
2.576 | -34.0971 | -4.8222
2.5821 | -34.0972 | -4.8336
2.5882 | -34.0973 | -4.8445
2.5943 | -34.0974 | -4.8553
2.6004 | -34.0976 | -4.8665
2.6065 | -34.0977 | -4.8771
2.6126 | -34.0978 | -4.8875
2.6187 | -34.0979 | -4.898
2.6248 | -34.098 | -4.9088
2.6309 | -34.0982 | -4.9204
2.637 | -34.0984 | -4.9325
2.6431 | -34.0986 | -4.943
2.6492 | -34.0987 | -4.9544
2.6554 | -34.0989 | -4.9657
2.6615 | -34.099 | -4.9769
2.6676 | -34.0992 | -4.9884
2.6737 | -34.0995 | -5
2.6798 | -34.0998 | -5.0112
2.6859 | -34.1 | -5.0228
2.692 | -34.1001 | -5.034
2.6981 | -34.1003 | -5.0454
2.7042 | -34.1005 | -5.0569
2.7103 | -34.1007 | -5.0679
2.7164 | -34.101 | -5.0792
2.7225 | -34.1013 | -5.0907
2.7286 | -34.1015 | -5.1021
2.7347 | -34.1017 | -5.1129
2.7408 | -34.102 | -5.1234
2.7469 | -34.1022 | -5.1349
2.753 | -34.1025 | -5.1458
2.7591 | -34.1028 | -5.1574
2.7652 | -34.1031 | -5.1693
2.7713 | -34.1032 | -5.1804
2.7774 | -34.1035 | -5.1918
2.7835 | -34.1037 | -5.2031
2.7896 | -34.104 | -5.2139
2.7958 | -34.1042 | -5.2252
2.8019 | -34.1045 | -5.2361
2.808 | -34.1048 | -5.2472
2.8141 | -34.1052 | -5.2587
2.8202 | -34.1055 | -5.2699
2.8263 | -34.1058 | -5.2812
2.8324 | -34.1062 | -5.2931
2.8385 | -34.1066 | -5.3045
2.8446 | -34.107 | -5.3156
2.8507 | -34.1074 | -5.3269
2.8568 | -34.1078 | -5.3378
2.8629 | -34.1081 | -5.3487
2.869 | -34.1085 | -5.3602
2.8751 | -34.1088 | -5.3713
2.8812 | -34.1093 | -5.3831
2.8873 | -34.1096 | -5.3942
2.8934 | -34.1099 | -5.4055
2.8995 | -34.1103 | -5.4168
2.9056 | -34.1106 | -5.4283
2.9117 | -34.111 | -5.4395
2.9178 | -34.1114 | -5.4504
2.9239 | -34.1118 | -5.4622
2.93 | -34.1123 | -5.474
2.9361 | -34.1126 | -5.4851
2.9423 | -34.113 | -5.4961
2.9484 | -34.1134 | -5.5073
2.9545 | -34.1139 | -5.5185
2.9606 | -34.1143 | -5.5302
2.9667 | -34.1147 | -5.5417
2.9728 | -34.1151 | -5.5531
2.9789 | -34.1156 | -5.5646
2.985 | -34.1161 | -5.5762
2.9911 | -34.1165 | -5.5877
2.9972 | -34.117 | -5.599
3.0033 | -34.1174 | -5.6107
3.0094 | -34.1178 | -5.6219
3.0155 | -34.1183 | -5.633
3.0216 | -34.1187 | -5.6444
3.0277 | -34.1191 | -5.6563
3.0338 | -34.1195 | -5.6681
3.0399 | -34.12 | -5.6797
3.046 | -34.1203 | -5.6913
3.0521 | -34.1207 | -5.7028
3.0582 | -34.121 | -5.7151
3.0643 | -34.1213 | -5.7269
3.0704 | -34.1217 | -5.739
3.0765 | -34.122 | -5.7503
3.0827 | -34.1224 | -5.7619
3.0888 | -34.1227 | -5.7734
3.0949 | -34.1231 | -5.7847
3.101 | -34.1235 | -5.7964
3.1071 | -34.1238 | -5.8084
3.1132 | -34.1242 | -5.8205
3.1193 | -34.1246 | -5.8322
3.1254 | -34.125 | -5.8438
3.1315 | -34.1253 | -5.8558
3.1376 | -34.1256 | -5.8675
3.1437 | -34.126 | -5.8792
3.1498 | -34.1264 | -5.8915
3.1559 | -34.1267 | -5.9031
3.162 | -34.1269 | -5.9154
3.1681 | -34.1272 | -5.9275
3.1742 | -34.1275 | -5.9393
3.1803 | -34.1278 | -5.9512
3.1864 | -34.1282 | -5.9627
3.1925 | -34.1285 | -5.9744
3.1986 | -34.1288 | -5.9866
3.2047 | -34.129 | -5.9988
3.2108 | -34.1293 | -6.0101
3.2169 | -34.1296 | -6.0219
3.223 | -34.1299 | -6.0339
3.2292 | -34.1301 | -6.0459
3.2353 | -34.1303 | -6.057
3.2414 | -34.1305 | -6.0686
3.2475 | -34.1307 | -6.0801
3.2536 | -34.131 | -6.0919
3.2597 | -34.1313 | -6.1031
3.2658 | -34.1315 | -6.1142
3.2719 | -34.1317 | -6.1254
3.278 | -34.1319 | -6.1371
3.2841 | -34.1321 | -6.1484
3.2902 | -34.1324 | -6.1601
3.2963 | -34.1326 | -6.1718
3.3024 | -34.1327 | -6.1831
3.3085 | -34.133 | -6.1949
3.3146 | -34.1331 | -6.2062
3.3207 | -34.1334 | -6.2178
3.3268 | -34.1336 | -6.2295
3.3329 | -34.1337 | -6.2405
3.339 | -34.1339 | -6.2514
3.3451 | -34.134 | -6.2628
3.3512 | -34.1342 | -6.2737
3.3573 | -34.1344 | -6.2852
3.3634 | -34.1346 | -6.2964
3.3696 | -34.1348 | -6.308
3.3757 | -34.1349 | -6.319
3.3818 | -34.1351 | -6.33
3.3879 | -34.1353 | -6.3414
3.394 | -34.1354 | -6.3528
3.4001 | -34.1356 | -6.3639
3.4062 | -34.1357 | -6.3751
3.4123 | -34.1358 | -6.3858
3.4184 | -34.136 | -6.3972
3.4245 | -34.1363 | -6.4081
3.4306 | -34.1363 | -6.4188
3.4367 | -34.1365 | -6.4295
3.4428 | -34.1366 | -6.44
3.4489 | -34.1368 | -6.4511
3.455 | -34.1369 | -6.4621
3.4611 | -34.1371 | -6.4734
3.4672 | -34.1372 | -6.4846
3.4733 | -34.1374 | -6.4957
3.4794 | -34.1375 | -6.5066
3.4855 | -34.1377 | -6.5171
3.4916 | -34.1379 | -6.5275
3.4977 | -34.138 | -6.539
3.5038 | -34.1382 | -6.55
3.5099 | -34.1383 | -6.5604
3.5161 | -34.1384 | -6.5708
3.5222 | -34.1386 | -6.5813
3.5283 | -34.1386 | -6.5922
3.5344 | -34.1388 | -6.603
3.5405 | -34.1389 | -6.6136
3.5466 | -34.1391 | -6.6247
3.5527 | -34.1392 | -6.6354
3.5588 | -34.1393 | -6.6462
3.5649 | -34.1394 | -6.6569
3.571 | -34.1395 | -6.6676
3.5771 | -34.1396 | -6.6788
3.5832 | -34.1398 | -6.6886
3.5893 | -34.14 | -6.6994
3.5954 | -34.1402 | -6.7102
3.6015 | -34.1403 | -6.7211
3.6076 | -34.1405 | -6.7316
3.6137 | -34.1407 | -6.7417
3.6198 | -34.141 | -6.7519
3.6259 | -34.1411 | -6.7629
3.632 | -34.1413 | -6.7735
3.6381 | -34.1415 | -6.7843
3.6442 | -34.1416 | -6.7948
3.6503 | -34.1418 | -6.8051
3.6565 | -34.1421 | -6.8158
3.6626 | -34.1422 | -6.8264
3.6687 | -34.1425 | -6.8372
3.6748 | -34.1427 | -6.8476
3.6809 | -34.143 | -6.8582
3.687 | -34.1431 | -6.8692
3.6931 | -34.1434 | -6.8794
3.6992 | -34.1437 | -6.89
3.7053 | -34.1439 | -6.9005
3.7114 | -34.144 | -6.9111
3.7175 | -34.1443 | -6.9217
3.7236 | -34.1446 | -6.9317
3.7297 | -34.1448 | -6.9424
3.7358 | -34.145 | -6.9533
3.7419 | -34.1452 | -6.9642
3.748 | -34.1455 | -6.9747
3.7541 | -34.1457 | -6.9852
3.7602 | -34.146 | -6.9956
3.7663 | -34.1462 | -7.006
3.7724 | -34.1465 | -7.0167
3.7785 | -34.1468 | -7.027
3.7846 | -34.147 | -7.0374
3.7907 | -34.1473 | -7.0482
3.7969 | -34.1476 | -7.059
3.803 | -34.1479 | -7.0692
3.8091 | -34.1482 | -7.0797
3.8152 | -34.1485 | -7.0901
3.8213 | -34.1487 | -7.101
3.8274 | -34.149 | -7.1112
3.8335 | -34.1494 | -7.1219
3.8396 | -34.1497 | -7.1326
3.8457 | -34.15 | -7.1433
3.8518 | -34.1504 | -7.1539
3.8579 | -34.1507 | -7.1645
3.864 | -34.151 | -7.1752
3.8701 | -34.1514 | -7.1861
3.8762 | -34.1516 | -7.1966
3.8823 | -34.1519 | -7.2074
3.8884 | -34.1523 | -7.218
3.8945 | -34.1527 | -7.2289
3.9006 | -34.153 | -7.2399
3.9067 | -34.1533 | -7.251
3.9128 | -34.1536 | -7.2617
3.9189 | -34.1539 | -7.2722
3.925 | -34.1542 | -7.2829
3.9311 | -34.1545 | -7.2932
3.9372 | -34.1549 | -7.304
3.9434 | -34.1552 | -7.3148
3.9495 | -34.1556 | -7.3256
3.9556 | -34.156 | -7.336
3.9617 | -34.1563 | -7.3463
3.9678 | -34.1567 | -7.3566
3.9739 | -34.1569 | -7.3673
3.98 | -34.1572 | -7.3781
3.9861 | -34.1576 | -7.3887
3.9922 | -34.1578 | -7.3994
3.9983 | -34.1582 | -7.4097
4.0044 | -34.1585 | -7.4199
4.0105 | -34.1588 | -7.4306
4.0166 | -34.1592 | -7.4413
4.0227 | -34.1595 | -7.4519
4.0288 | -34.1598 | -7.4631
4.0349 | -34.16 | -7.4735
4.041 | -34.1603 | -7.4841
4.0471 | -34.1606 | -7.4948
4.0532 | -34.1609 | -7.5052
4.0593 | -34.1613 | -7.516
4.0654 | -34.1616 | -7.5265
4.0715 | -34.1619 | -7.5376
4.0776 | -34.1623 | -7.5482
4.0838 | -34.1626 | -7.5592
4.0899 | -34.163 | -7.5696
4.096 | -34.1633 | -7.5802
4.1021 | -34.1636 | -7.5906
4.1082 | -34.1639 | -7.601
4.1143 | -34.1641 | -7.6114
4.1204 | -34.1645 | -7.6217
4.1265 | -34.1648 | -7.6321
4.1326 | -34.1651 | -7.6424
4.1387 | -34.1655 | -7.6528
4.1448 | -34.1658 | -7.6635
4.1509 | -34.166 | -7.6741
4.157 | -34.1663 | -7.6852
4.1631 | -34.1666 | -7.696
4.1692 | -34.167 | -7.7064
4.1753 | -34.1673 | -7.7169
4.1814 | -34.1675 | -7.7271
4.1875 | -34.1678 | -7.738
4.1936 | -34.1681 | -7.7487
4.1997 | -34.1684 | -7.759
4.2058 | -34.1687 | -7.7698
4.2119 | -34.169 | -7.7802
4.218 | -34.1692 | -7.7905
4.2241 | -34.1695 | -7.8008
4.2303 | -34.1698 | -7.8112
4.2364 | -34.1701 | -7.8217
4.2425 | -34.1703 | -7.8322
4.2486 | -34.1706 | -7.8428
4.2547 | -34.1707 | -7.8531
4.2608 | -34.171 | -7.8637
4.2669 | -34.1712 | -7.8739
4.273 | -34.1715 | -7.8846
4.2791 | -34.1718 | -7.8951
4.2852 | -34.172 | -7.9057
4.2913 | -34.1722 | -7.9159
4.2974 | -34.1725 | -7.9267
4.3035 | -34.1728 | -7.9375
4.3096 | -34.173 | -7.9481
4.3157 | -34.1733 | -7.9584
4.3218 | -34.1735 | -7.9688
4.3279 | -34.1738 | -7.9793
4.334 | -34.174 | -7.9896
4.3401 | -34.1743 | -7.9997
4.3462 | -34.1746 | -8.0102
4.3523 | -34.1749 | -8.0202
4.3584 | -34.175 | -8.0302
4.3645 | -34.1753 | -8.0405
4.3707 | -34.1755 | -8.0506
4.3768 | -34.1758 | -8.0609
4.3829 | -34.176 | -8.0718
4.389 | -34.1763 | -8.0821
4.3951 | -34.1765 | -8.0923
4.4012 | -34.1767 | -8.103
4.4073 | -34.1769 | -8.1132
4.4134 | -34.1772 | -8.1235
4.4195 | -34.1775 | -8.1338
4.4256 | -34.1778 | -8.144
4.4317 | -34.1781 | -8.1542
4.4378 | -34.1783 | -8.1642
4.4439 | -34.1785 | -8.1742
4.45 | -34.1787 | -8.1845
4.4561 | -34.1789 | -8.1948
4.4622 | -34.1791 | -8.2052
4.4683 | -34.1793 | -8.2155
4.4744 | -34.1796 | -8.2261
4.4805 | -34.1798 | -8.236
4.4866 | -34.18 | -8.2463
4.4927 | -34.1803 | -8.2569
4.4988 | -34.1805 | -8.2671
4.5049 | -34.1808 | -8.277
4.511 | -34.1811 | -8.287
4.5172 | -34.1813 | -8.2972
4.5233 | -34.1815 | -8.3079
4.5294 | -34.1817 | -8.318
4.5355 | -34.182 | -8.3283
4.5416 | -34.1822 | -8.3385
4.5477 | -34.1825 | -8.3488
4.5538 | -34.1826 | -8.3586
4.5599 | -34.1829 | -8.3688
4.566 | -34.1832 | -8.379
4.5721 | -34.1834 | -8.3893
4.5782 | -34.1837 | -8.3998
4.5843 | -34.1839 | -8.4102
4.5904 | -34.1842 | -8.4205
4.5965 | -34.1845 | -8.4306
4.6026 | -34.1847 | -8.4414
4.6087 | -34.1849 | -8.4515
4.6148 | -34.1851 | -8.4621
4.6209 | -34.1854 | -8.4723
4.627 | -34.1857 | -8.4828
4.6331 | -34.1858 | -8.4929
4.6392 | -34.1861 | -8.5033
4.6453 | -34.1863 | -8.5135
4.6514 | -34.1866 | -8.5238
4.6576 | -34.1868 | -8.5345
4.6637 | -34.187 | -8.5444
4.6698 | -34.1872 | -8.5545
4.6759 | -34.1875 | -8.5651
4.682 | -34.1877 | -8.5753
4.6881 | -34.1878 | -8.5856
4.6942 | -34.1881 | -8.5958
4.7003 | -34.1883 | -8.6061
4.7064 | -34.1886 | -8.6162
4.7125 | -34.1889 | -8.6268
4.7186 | -34.1891 | -8.6367
4.7247 | -34.1893 | -8.6471
4.7308 | -34.1896 | -8.6574
4.7369 | -34.1898 | -8.6678
4.743 | -34.1901 | -8.6784
4.7491 | -34.1903 | -8.689
4.7552 | -34.1905 | -8.6995
4.7613 | -34.1907 | -8.7099
4.7674 | -34.1909 | -8.7201
4.7735 | -34.1912 | -8.7308
4.7796 | -34.1915 | -8.7414
4.7857 | -34.1917 | -8.7518
4.7918 | -34.192 | -8.762
4.7979 | -34.1922 | -8.7722
4.8041 | -34.1925 | -8.7826
4.8102 | -34.1927 | -8.7926
4.8163 | -34.193 | -8.8029
4.8224 | -34.1932 | -8.8133
4.8285 | -34.1935 | -8.8232
4.8346 | -34.1938 | -8.8334
4.8407 | -34.1939 | -8.8437
4.8468 | -34.1942 | -8.8544
4.8529 | -34.1945 | -8.865
4.859 | -34.1947 | -8.8756
4.8651 | -34.195 | -8.8858
4.8712 | -34.1952 | -8.8959
4.8773 | -34.1954 | -8.9064
4.8834 | -34.1957 | -8.9168
4.8895 | -34.196 | -8.9269
4.8956 | -34.1962 | -8.9372
4.9017 | -34.1964 | -8.9473
4.9078 | -34.1966 | -8.9579
4.9139 | -34.1969 | -8.9679
4.92 | -34.1972 | -8.9779
4.9261 | -34.1975 | -8.9879
4.9322 | -34.1977 | -8.9983
4.9383 | -34.198 | -9.0086
4.9445 | -34.1982 | -9.0189
4.9506 | -34.1985 | -9.0291
4.9567 | -34.1987 | -9.0392
4.9628 | -34.199 | -9.0494
4.9689 | -34.1992 | -9.0602
4.975 | -34.1995 | -9.0708
4.9811 | -34.1997 | -9.0808
4.9872 | -34.2001 | -9.0911
4.9933 | -34.2004 | -9.1017
4.9994 | -34.2006 | -9.1124
5.0055 | -34.2009 | -9.1228
5.0116 | -34.2012 | -9.1336
5.0177 | -34.2015 | -9.1442
5.0238 | -34.2018 | -9.1549
5.0299 | -34.2022 | -9.1656
5.036 | -34.2024 | -9.176
5.0421 | -34.2027 | -9.1863
5.0482 | -34.203 | -9.1967
5.0543 | -34.2032 | -9.2068
5.0604 | -34.2035 | -9.2172
5.0665 | -34.2038 | -9.2273
5.0726 | -34.2041 | -9.238
5.0787 | -34.2044 | -9.2486
5.0848 | -34.2047 | -9.2591
5.091 | -34.205 | -9.2693
5.0971 | -34.2053 | -9.2794
5.1032 | -34.2056 | -9.2901
5.1093 | -34.2059 | -9.3005
5.1154 | -34.2061 | -9.3107
5.1215 | -34.2064 | -9.3208
5.1276 | -34.2067 | -9.331
5.1337 | -34.207 | -9.3411
5.1398 | -34.2073 | -9.3513
5.1459 | -34.2076 | -9.3619
5.152 | -34.2079 | -9.3725
5.1581 | -34.2081 | -9.3831
5.1642 | -34.2084 | -9.3936
5.1703 | -34.2087 | -9.404
5.1764 | -34.2091 | -9.4142
5.1825 | -34.2093 | -9.4243
5.1886 | -34.2097 | -9.4344
5.1947 | -34.2099 | -9.4449
5.2008 | -34.2102 | -9.4554
5.2069 | -34.2105 | -9.466
5.213 | -34.2108 | -9.4762
5.2191 | -34.2111 | -9.4867
5.2252 | -34.2114 | -9.4971
5.2314 | -34.2117 | -9.5077
5.2375 | -34.212 | -9.5179
5.2436 | -34.2123 | -9.5283
5.2497 | -34.2126 | -9.5385
5.2558 | -34.2129 | -9.5493
5.2619 | -34.2132 | -9.56
5.268 | -34.2135 | -9.5705
5.2741 | -34.2139 | -9.5811
5.2802 | -34.2141 | -9.5919
5.2863 | -34.2145 | -9.6022
5.2924 | -34.2148 | -9.6129
5.2985 | -34.2151 | -9.6234
5.3046 | -34.2154 | -9.6339
5.3107 | -34.2157 | -9.6446
5.3168 | -34.2161 | -9.6552
5.3229 | -34.2164 | -9.6657
5.329 | -34.2167 | -9.6761
5.3351 | -34.217 | -9.6863
5.3412 | -34.2173 | -9.6968
5.3473 | -34.2176 | -9.7073
5.3534 | -34.2179 | -9.7179
5.3595 | -34.2182 | -9.7285
5.3656 | -34.2186 | -9.7389
5.3717 | -34.2189 | -9.7498
5.3779 | -34.2192 | -9.7606
5.384 | -34.2195 | -9.7711
5.3901 | -34.2198 | -9.7818
5.3962 | -34.2201 | -9.7921
5.4023 | -34.2205 | -9.8029
5.4084 | -34.2208 | -9.8133
5.4145 | -34.2212 | -9.8236
5.4206 | -34.2215 | -9.8342
5.4267 | -34.2218 | -9.8451
5.4328 | -34.2221 | -9.8556
5.4389 | -34.2224 | -9.8659
5.445 | -34.2227 | -9.8768
5.4511 | -34.2231 | -9.8871
5.4572 | -34.2234 | -9.8977
5.4633 | -34.2237 | -9.9082
5.4694 | -34.224 | -9.9188
5.4755 | -34.2243 | -9.929
5.4816 | -34.2246 | -9.9396
5.4877 | -34.225 | -9.9501
5.4938 | -34.2253 | -9.9603
5.4999 | -34.2256 | -9.9708
5.506 | -34.226 | -9.9813
5.5121 | -34.2263 | -9.9917
5.5183 | -34.2266 | -10.0021
5.5244 | -34.2269 | -10.0128
5.5305 | -34.2273 | -10.0233
5.5366 | -34.2276 | -10.0335
5.5427 | -34.2279 | -10.0442
5.5488 | -34.2282 | -10.0548
5.5549 | -34.2285 | -10.0654
5.561 | -34.2288 | -10.0762
5.5671 | -34.2292 | -10.0865
5.5732 | -34.2295 | -10.0973
5.5793 | -34.2298 | -10.1083
5.5854 | -34.2302 | -10.1187
5.5915 | -34.2305 | -10.1296
5.5976 | -34.2308 | -10.1402
5.6037 | -34.2311 | -10.1508
5.6098 | -34.2314 | -10.1613
5.6159 | -34.2318 | -10.1717
5.622 | -34.2321 | -10.1823
5.6281 | -34.2324 | -10.193
5.6342 | -34.2328 | -10.2039
5.6403 | -34.2331 | -10.2145
5.6464 | -34.2334 | -10.2252
5.6525 | -34.2337 | -10.2358
5.6586 | -34.2341 | -10.2461
5.6648 | -34.2344 | -10.2567
5.6709 | -34.2348 | -10.2675
5.677 | -34.2351 | -10.2781
5.6831 | -34.2355 | -10.2887
5.6892 | -34.2358 | -10.2995
5.6953 | -34.2361 | -10.3101
5.7014 | -34.2364 | -10.3208
5.7075 | -34.2367 | -10.3313
5.7136 | -34.237 | -10.3419
5.7197 | -34.2373 | -10.3524
5.7258 | -34.2377 | -10.3628
5.7319 | -34.238 | -10.3734
5.738 | -34.2384 | -10.384
5.7441 | -34.2387 | -10.3945
5.7502 | -34.2391 | -10.4048
5.7563 | -34.2394 | -10.4154
5.7624 | -34.2398 | -10.4262
5.7685 | -34.2401 | -10.4368
5.7746 | -34.2404 | -10.4477
5.7807 | -34.2407 | -10.4583
5.7868 | -34.241 | -10.4691
5.7929 | -34.2413 | -10.4795
5.799 | -34.2416 | -10.49
5.8052 | -34.2419 | -10.5004
5.8113 | -34.2423 | -10.5109
5.8174 | -34.2426 | -10.5214
5.8235 | -34.2429 | -10.5319
5.8296 | -34.2432 | -10.5426
5.8357 | -34.2435 | -10.5532
5.8418 | -34.2439 | -10.5639
5.8479 | -34.2441 | -10.5745
5.854 | -34.2445 | -10.5852
5.8601 | -34.2448 | -10.5956
5.8662 | -34.2451 | -10.6064
5.8723 | -34.2454 | -10.6168
5.8784 | -34.2457 | -10.6273
5.8845 | -34.2461 | -10.6377
5.8906 | -34.2464 | -10.6483
5.8967 | -34.2467 | -10.6587
5.9028 | -34.2471 | -10.6691
5.9089 | -34.2474 | -10.6795
5.915 | -34.2477 | -10.6898
5.9211 | -34.248 | -10.7003
5.9272 | -34.2484 | -10.7106
5.9333 | -34.2487 | -10.721
5.9394 | -34.249 | -10.7315
5.9455 | -34.2494 | -10.7418
5.9517 | -34.2498 | -10.7522
5.9578 | -34.2501 | -10.7628
5.9639 | -34.2505 | -10.773
5.97 | -34.2508 | -10.7836
5.9761 | -34.2512 | -10.794
5.9822 | -34.2515 | -10.8045
5.9883 | -34.2518 | -10.815
5.9944 | -34.2522 | -10.8256
6.0005 | -34.2525 | -10.8363
6.0066 | -34.2528 | -10.8465
6.0127 | -34.2532 | -10.8568
6.0188 | -34.2535 | -10.8675
6.0249 | -34.2538 | -10.8781
6.031 | -34.2542 | -10.8884
6.0371 | -34.2545 | -10.8989
6.0432 | -34.2548 | -10.9093
6.0493 | -34.2552 | -10.9196
6.0554 | -34.2555 | -10.9299
6.0615 | -34.2559 | -10.9403
6.0676 | -34.2562 | -10.951
6.0737 | -34.2566 | -10.9616
6.0798 | -34.2569 | -10.972
6.0859 | -34.2573 | -10.9824
6.0921 | -34.2576 | -10.9926
6.0982 | -34.2579 | -11.003
6.1043 | -34.2583 | -11.0133
6.1104 | -34.2586 | -11.024
6.1165 | -34.259 | -11.0339
6.1226 | -34.2593 | -11.0443
6.1287 | -34.2597 | -11.0547
6.1348 | -34.26 | -11.0651
6.1409 | -34.2603 | -11.0754
6.147 | -34.2607 | -11.0857
6.1531 | -34.261 | -11.0961
6.1592 | -34.2613 | -11.1063
6.1653 | -34.2616 | -11.1166
6.1714 | -34.262 | -11.127
6.1775 | -34.2623 | -11.1375
6.1836 | -34.2626 | -11.148
6.1897 | -34.263 | -11.1583
6.1958 | -34.2633 | -11.169
6.2019 | -34.2637 | -11.1795
6.208 | -34.2641 | -11.1899
6.2141 | -34.2644 | -11.2002
6.2202 | -34.2648 | -11.2107
6.2263 | -34.2651 | -11.2209
6.2325 | -34.2654 | -11.2309
6.2386 | -34.2658 | -11.241
6.2447 | -34.2661 | -11.2516
6.2508 | -34.2665 | -11.262
6.2569 | -34.2669 | -11.2722
6.263 | -34.2672 | -11.2826
6.2691 | -34.2675 | -11.2928
6.2752 | -34.2678 | -11.3033
6.2813 | -34.2681 | -11.3136
6.2874 | -34.2685 | -11.3238
6.2935 | -34.2688 | -11.3341
6.2996 | -34.2691 | -11.3444
6.3057 | -34.2694 | -11.3548
6.3118 | -34.2697 | -11.3652
6.3179 | -34.2701 | -11.3755
6.324 | -34.2704 | -11.3857
6.3301 | -34.2708 | -11.396
6.3362 | -34.2711 | -11.4062
6.3423 | -34.2715 | -11.4167
6.3484 | -34.2718 | -11.4271
6.3545 | -34.2721 | -11.437
6.3606 | -34.2725 | -11.447
6.3667 | -34.2727 | -11.4571
6.3728 | -34.2731 | -11.4675
6.379 | -34.2734 | -11.4778
6.3851 | -34.2738 | -11.4878
6.3912 | -34.2741 | -11.4981
6.3973 | -34.2744 | -11.5084
6.4034 | -34.2747 | -11.5184
6.4095 | -34.2751 | -11.5287
6.4156 | -34.2754 | -11.539
6.4217 | -34.2757 | -11.5492
6.4278 | -34.276 | -11.5595
6.4339 | -34.2764 | -11.5697
6.44 | -34.2768 | -11.5798
6.4461 | -34.2771 | -11.5903
6.4522 | -34.2774 | -11.6004
6.4583 | -34.2777 | -11.6108
6.4644 | -34.278 | -11.6211
6.4705 | -34.2783 | -11.631
6.4766 | -34.2787 | -11.6413
6.4827 | -34.279 | -11.6516
6.4888 | -34.2793 | -11.6616
6.4949 | -34.2796 | -11.6716
6.501 | -34.28 | -11.6819
6.5071 | -34.2803 | -11.6918
6.5132 | -34.2807 | -11.7021
6.5194 | -34.281 | -11.7123
6.5255 | -34.2813 | -11.7223
6.5316 | -34.2816 | -11.7327
6.5377 | -34.2819 | -11.7428
6.5438 | -34.2822 | -11.7528
6.5499 | -34.2826 | -11.7629
6.556 | -34.2829 | -11.773
6.5621 | -34.2832 | -11.7835
6.5682 | -34.2835 | -11.7937
6.5743 | -34.2839 | -11.8038
6.5804 | -34.2842 | -11.8141
6.5865 | -34.2846 | -11.8242
6.5926 | -34.2849 | -11.8343
6.5987 | -34.2852 | -11.8441
6.6048 | -34.2856 | -11.8543
6.6109 | -34.2859 | -11.8645
6.617 | -34.2862 | -11.8748
6.6231 | -34.2865 | -11.8849
6.6292 | -34.2869 | -11.895
6.6353 | -34.2872 | -11.9055
6.6414 | -34.2875 | -11.9155
6.6475 | -34.2879 | -11.9257
6.6536 | -34.2882 | -11.9357
6.6597 | -34.2885 | -11.9455
6.6659 | -34.2888 | -11.9557
6.672 | -34.2892 | -11.9657
6.6781 | -34.2895 | -11.9759
6.6842 | -34.2898 | -11.9858
6.6903 | -34.2902 | -11.9959
6.6964 | -34.2906 | -12.006
6.7025 | -34.2909 | -12.0159
6.7086 | -34.2912 | -12.026
6.7147 | -34.2915 | -12.0362
6.7208 | -34.2918 | -12.0464
6.7269 | -34.2921 | -12.0564
6.733 | -34.2924 | -12.0663
6.7391 | -34.2928 | -12.0764
6.7452 | -34.2931 | -12.0862
6.7513 | -34.2935 | -12.096
6.7574 | -34.2938 | -12.1062
6.7635 | -34.2941 | -12.116
6.7696 | -34.2945 | -12.126
6.7757 | -34.2948 | -12.1361
6.7818 | -34.2952 | -12.1461
6.7879 | -34.2955 | -12.156
6.794 | -34.2959 | -12.1658
6.8001 | -34.2961 | -12.1758
6.8063 | -34.2965 | -12.1858
6.8124 | -34.2968 | -12.196
6.8185 | -34.2971 | -12.2062
6.8246 | -34.2974 | -12.216
6.8307 | -34.2977 | -12.2259
6.8368 | -34.2981 | -12.2362
6.8429 | -34.2984 | -12.2463
6.849 | -34.2988 | -12.2563
6.8551 | -34.2991 | -12.2662
6.8612 | -34.2994 | -12.2762
6.8673 | -34.2997 | -12.2862
6.8734 | -34.3001 | -12.2963
6.8795 | -34.3004 | -12.3063
6.8856 | -34.3008 | -12.3164
6.8917 | -34.3011 | -12.3263
6.8978 | -34.3014 | -12.336
6.9039 | -34.3017 | -12.3457
6.91 | -34.3021 | -12.3557
6.9161 | -34.3024 | -12.3655
6.9222 | -34.3027 | -12.3755
6.9283 | -34.3031 | -12.3851
6.9344 | -34.3035 | -12.3951
6.9405 | -34.3038 | -12.405
6.9466 | -34.3042 | -12.4152
6.9528 | -34.3045 | -12.4253
6.9589 | -34.3049 | -12.4354
6.965 | -34.3052 | -12.4455
6.9711 | -34.3056 | -12.4555
6.9772 | -34.306 | -12.4655
6.9833 | -34.3063 | -12.4754
6.9894 | -34.3066 | -12.4856
6.9955 | -34.307 | -12.4954
7.0016 | -34.3073 | -12.5053
7.0077 | -34.3076 | -12.5152
7.0138 | -34.308 | -12.5251
7.0199 | -34.3083 | -12.5349
7.026 | -34.3087 | -12.5448
7.0321 | -34.309 | -12.5545
7.0382 | -34.3094 | -12.5641
7.0443 | -34.3097 | -12.5743
7.0504 | -34.3101 | -12.5843
7.0565 | -34.3104 | -12.5942
7.0626 | -34.3108 | -12.6039
7.0687 | -34.3111 | -12.6133
7.0748 | -34.3115 | -12.6229
7.0809 | -34.3118 | -12.6328
7.087 | -34.3121 | -12.6428
7.0932 | -34.3124 | -12.6527
7.0993 | -34.3128 | -12.6625
7.1054 | -34.3131 | -12.6722
7.1115 | -34.3135 | -12.6818
7.1176 | -34.3138 | -12.6917
7.1237 | -34.3141 | -12.7015
7.1298 | -34.3145 | -12.7113
7.1359 | -34.3148 | -12.7212
7.142 | -34.3152 | -12.7307
7.1481 | -34.3155 | -12.7405
7.1542 | -34.3159 | -12.7503
7.1603 | -34.3162 | -12.7601
7.1664 | -34.3166 | -12.7699
7.1725 | -34.317 | -12.7798
7.1786 | -34.3174 | -12.7896
7.1847 | -34.3177 | -12.7994
7.1908 | -34.318 | -12.8094
7.1969 | -34.3183 | -12.8194
7.203 | -34.3187 | -12.8291
7.2091 | -34.319 | -12.8388
7.2152 | -34.3194 | -12.8486
7.2213 | -34.3197 | -12.8582
7.2274 | -34.3201 | -12.868
7.2335 | -34.3204 | -12.8778
7.2397 | -34.3208 | -12.8876
7.2458 | -34.3212 | -12.8972
7.2519 | -34.3215 | -12.907
7.258 | -34.3218 | -12.9169
7.2641 | -34.3222 | -12.9268
7.2702 | -34.3225 | -12.9367
7.2763 | -34.3229 | -12.9466
7.2824 | -34.3233 | -12.9566
7.2885 | -34.3236 | -12.9664
7.2946 | -34.324 | -12.9762
7.3007 | -34.3243 | -12.9861
7.3068 | -34.3246 | -12.9962
7.3129 | -34.325 | -13.0062
7.319 | -34.3254 | -13.0161
7.3251 | -34.3257 | -13.0256
7.3312 | -34.3261 | -13.0354
7.3373 | -34.3264 | -13.045
7.3434 | -34.3268 | -13.0547
7.3495 | -34.3272 | -13.0644
7.3556 | -34.3275 | -13.0743
7.3617 | -34.3279 | -13.0843
7.3678 | -34.3282 | -13.094
7.3739 | -34.3286 | -13.104
7.3801 | -34.3289 | -13.1141
7.3862 | -34.3293 | -13.1239
7.3923 | -34.3297 | -13.1338
7.3984 | -34.33 | -13.1434
7.4045 | -34.3304 | -13.1531
7.4106 | -34.3308 | -13.1629
7.4167 | -34.3311 | -13.1727
7.4228 | -34.3315 | -13.1827
7.4289 | -34.3319 | -13.1925
7.435 | -34.3322 | -13.202
7.4411 | -34.3326 | -13.2119
7.4472 | -34.3329 | -13.2218
7.4533 | -34.3333 | -13.2315
7.4594 | -34.3337 | -13.2415
7.4655 | -34.334 | -13.2511
7.4716 | -34.3344 | -13.2611
7.4777 | -34.3347 | -13.2709
7.4838 | -34.3351 | -13.2806
7.4899 | -34.3354 | -13.2904
7.496 | -34.3358 | -13.3
7.5021 | -34.3362 | -13.3098
7.5082 | -34.3365 | -13.3196
7.5143 | -34.3369 | -13.3294
7.5204 | -34.3372 | -13.3394
7.5266 | -34.3376 | -13.3491
7.5327 | -34.338 | -13.3588
7.5388 | -34.3383 | -13.3689
7.5449 | -34.3387 | -13.3789
7.551 | -34.3391 | -13.3885
7.5571 | -34.3394 | -13.3985
7.5632 | -34.3398 | -13.4084
7.5693 | -34.3402 | -13.4181
7.5754 | -34.3406 | -13.4278
7.5815 | -34.3409 | -13.4376
7.5876 | -34.3413 | -13.4476
7.5937 | -34.3417 | -13.4576
7.5998 | -34.3421 | -13.4673
7.6059 | -34.3425 | -13.4774
7.612 | -34.3428 | -13.4871
7.6181 | -34.3432 | -13.4969
7.6242 | -34.3435 | -13.507
7.6303 | -34.3439 | -13.5171
7.6364 | -34.3443 | -13.527
7.6425 | -34.3446 | -13.5369
7.6486 | -34.345 | -13.5468
7.6547 | -34.3454 | -13.5565
7.6608 | -34.3457 | -13.5663
7.667 | -34.3461 | -13.5762
7.6731 | -34.3465 | -13.5862
7.6792 | -34.3468 | -13.5962
7.6853 | -34.3472 | -13.6062
7.6914 | -34.3476 | -13.6161
7.6975 | -34.348 | -13.6261
7.7036 | -34.3484 | -13.6361
7.7097 | -34.3487 | -13.6461
7.7158 | -34.3491 | -13.6561
7.7219 | -34.3495 | -13.6658
7.728 | -34.3499 | -13.6758
7.7341 | -34.3502 | -13.6858
7.7402 | -34.3506 | -13.6958
7.7463 | -34.351 | -13.7057
7.7524 | -34.3513 | -13.7159
7.7585 | -34.3517 | -13.7258
7.7646 | -34.3521 | -13.7357
7.7707 | -34.3524 | -13.7459
7.7768 | -34.3528 | -13.7559
7.7829 | -34.3532 | -13.7661
7.789 | -34.3535 | -13.7763
7.7951 | -34.3539 | -13.7862
7.8012 | -34.3542 | -13.7961
7.8073 | -34.3547 | -13.8063
7.8135 | -34.3551 | -13.8163
7.8196 | -34.3555 | -13.8265
7.8257 | -34.3558 | -13.8366
7.8318 | -34.3562 | -13.8466
7.8379 | -34.3566 | -13.8565
7.844 | -34.3569 | -13.8664
7.8501 | -34.3573 | -13.8765
7.8562 | -34.3577 | -13.8867
7.8623 | -34.3581 | -13.8966
7.8684 | -34.3584 | -13.9067
7.8745 | -34.3588 | -13.9166
7.8806 | -34.3592 | -13.9268
7.8867 | -34.3596 | -13.9365
7.8928 | -34.36 | -13.9463
7.8989 | -34.3603 | -13.9563
7.905 | -34.3607 | -13.9664
7.9111 | -34.361 | -13.9765
7.9172 | -34.3614 | -13.9865
7.9233 | -34.3618 | -13.9965
7.9294 | -34.3622 | -14.0066
7.9355 | -34.3625 | -14.0164
7.9416 | -34.3629 | -14.0264
7.9477 | -34.3633 | -14.0363
7.9539 | -34.3636 | -14.0463
7.96 | -34.364 | -14.0564
7.9661 | -34.3644 | -14.0665
7.9722 | -34.3648 | -14.0764
7.9783 | -34.3652 | -14.0866
7.9844 | -34.3656 | -14.0966
7.9905 | -34.366 | -14.1068
7.9966 | -34.3663 | -14.117
8.0027 | -34.3667 | -14.127
8.0088 | -34.367 | -14.137
8.0149 | -34.3674 | -14.1471
8.021 | -34.3678 | -14.1572
8.0271 | -34.3682 | -14.1675
8.0332 | -34.3685 | -14.1775
8.0393 | -34.3689 | -14.1876
8.0454 | -34.3693 | -14.198
8.0515 | -34.3697 | -14.208
8.0576 | -34.3701 | -14.2183
8.0637 | -34.3705 | -14.2283
8.0698 | -34.3708 | -14.2382
8.0759 | -34.3712 | -14.2482
8.082 | -34.3716 | -14.2581
8.0881 | -34.3719 | -14.2683
8.0942 | -34.3723 | -14.2784
8.1004 | -34.3728 | -14.2886
8.1065 | -34.3731 | -14.2984
8.1126 | -34.3735 | -14.3085
8.1187 | -34.3739 | -14.3186
8.1248 | -34.3742 | -14.3286
8.1309 | -34.3746 | -14.3385
8.137 | -34.375 | -14.3484
8.1431 | -34.3754 | -14.3587
8.1492 | -34.3757 | -14.3687
8.1553 | -34.3761 | -14.3789
8.1614 | -34.3765 | -14.3892
8.1675 | -34.3769 | -14.3992
8.1736 | -34.3772 | -14.4094
8.1797 | -34.3776 | -14.4195
8.1858 | -34.378 | -14.4297
8.1919 | -34.3783 | -14.4397
8.198 | -34.3787 | -14.4495
8.2041 | -34.3791 | -14.4597
8.2102 | -34.3795 | -14.4697
8.2163 | -34.3799 | -14.4796
8.2224 | -34.3802 | -14.4896
8.2285 | -34.3806 | -14.4995
8.2346 | -34.381 | -14.5095
8.2408 | -34.3814 | -14.5194
8.2469 | -34.3818 | -14.5295
8.253 | -34.3822 | -14.5394
8.2591 | -34.3826 | -14.5496
8.2652 | -34.383 | -14.5597
8.2713 | -34.3834 | -14.5698
8.2774 | -34.3838 | -14.58
8.2835 | -34.3841 | -14.5903
8.2896 | -34.3845 | -14.6002
8.2957 | -34.3849 | -14.6103
8.3018 | -34.3853 | -14.6204
8.3079 | -34.3857 | -14.6303
8.314 | -34.3861 | -14.6404
8.3201 | -34.3865 | -14.6503
8.3262 | -34.3868 | -14.6604
8.3323 | -34.3872 | -14.6704
8.3384 | -34.3876 | -14.6806
8.3445 | -34.388 | -14.6906
8.3506 | -34.3884 | -14.7007
8.3567 | -34.3888 | -14.7109
8.3628 | -34.3891 | -14.7209
8.3689 | -34.3895 | -14.7309
8.375 | -34.3899 | -14.741
8.3812 | -34.3903 | -14.751
8.3873 | -34.3907 | -14.761
8.3934 | -34.3911 | -14.7712
8.3995 | -34.3914 | -14.7811
8.4056 | -34.3918 | -14.7911
8.4117 | -34.3922 | -14.8012
8.4178 | -34.3926 | -14.8114
8.4239 | -34.3929 | -14.8213
8.43 | -34.3933 | -14.8314
8.4361 | -34.3936 | -14.8415
8.4422 | -34.394 | -14.8515
8.4483 | -34.3944 | -14.8615
8.4544 | -34.3948 | -14.8716
8.4605 | -34.3952 | -14.8815
8.4666 | -34.3956 | -14.8918
8.4727 | -34.3959 | -14.9018
8.4788 | -34.3963 | -14.9119
8.4849 | -34.3967 | -14.9218
8.491 | -34.3971 | -14.9319
8.4971 | -34.3975 | -14.9417
8.5032 | -34.3978 | -14.9517
8.5093 | -34.3982 | -14.9616
8.5154 | -34.3986 | -14.9717
8.5215 | -34.399 | -14.9816
8.5277 | -34.3994 | -14.9916
8.5338 | -34.3998 | -15.0016
8.5399 | -34.4001 | -15.0116
8.546 | -34.4006 | -15.0214
8.5521 | -34.4009 | -15.0315
8.5582 | -34.4013 | -15.0414
8.5643 | -34.4017 | -15.0512
8.5704 | -34.4021 | -15.0609
8.5765 | -34.4024 | -15.071
8.5826 | -34.4028 | -15.0807
8.5887 | -34.4032 | -15.0906
8.5948 | -34.4036 | -15.1006
8.6009 | -34.404 | -15.1104
8.607 | -34.4044 | -15.1205
8.6131 | -34.4048 | -15.1305
8.6192 | -34.4052 | -15.1402
8.6253 | -34.4056 | -15.1501
8.6314 | -34.4059 | -15.16
8.6375 | -34.4063 | -15.17
8.6436 | -34.4067 | -15.1799
8.6497 | -34.4071 | -15.1897
8.6558 | -34.4074 | -15.1996
8.6619 | -34.4078 | -15.2096
8.6681 | -34.4082 | -15.2197
8.6742 | -34.4086 | -15.2296
8.6803 | -34.4089 | -15.2395
8.6864 | -34.4094 | -15.2492
8.6925 | -34.4097 | -15.259
8.6986 | -34.4101 | -15.269
8.7047 | -34.4105 | -15.2791
8.7108 | -34.4108 | -15.289
8.7169 | -34.4113 | -15.299
8.723 | -34.4116 | -15.3087
8.7291 | -34.412 | -15.3184
8.7352 | -34.4124 | -15.3285
8.7413 | -34.4128 | -15.3385
8.7474 | -34.4131 | -15.3483
8.7535 | -34.4136 | -15.3581
8.7596 | -34.414 | -15.3681
8.7657 | -34.4143 | -15.3781
8.7718 | -34.4147 | -15.388
8.7779 | -34.4151 | -15.398
8.784 | -34.4155 | -15.4079
8.7901 | -34.4159 | -15.4177
8.7962 | -34.4163 | -15.4273
8.8023 | -34.4166 | -15.4372
8.8084 | -34.417 | -15.447
8.8146 | -34.4174 | -15.4569
8.8207 | -34.4178 | -15.4667
8.8268 | -34.4182 | -15.4762
8.8329 | -34.4186 | -15.4861
8.839 | -34.419 | -15.4959
8.8451 | -34.4193 | -15.5058
8.8512 | -34.4197 | -15.5155
8.8573 | -34.4201 | -15.5254
8.8634 | -34.4205 | -15.535
8.8695 | -34.4209 | -15.5451
8.8756 | -34.4212 | -15.5549
8.8817 | -34.4216 | -15.5646
8.8878 | -34.422 | -15.5747
8.8939 | -34.4224 | -15.5845
8.9 | -34.4228 | -15.5943
8.9061 | -34.4232 | -15.6041
8.9122 | -34.4235 | -15.614
8.9183 | -34.4239 | -15.6239
8.9244 | -34.4243 | -15.6338
8.9305 | -34.4248 | -15.6435
8.9366 | -34.4251 | -15.6533
8.9427 | -34.4255 | -15.6632
8.9488 | -34.4259 | -15.6729
8.955 | -34.4263 | -15.6827
8.9611 | -34.4267 | -15.6927
8.9672 | -34.4271 | -15.7025
8.9733 | -34.4275 | -15.7121
8.9794 | -34.4279 | -15.7218
8.9855 | -34.4283 | -15.7317
8.9916 | -34.4287 | -15.7417
8.9977 | -34.4291 | -15.7515
9.0038 | -34.4295 | -15.7611
9.0099 | -34.4299 | -15.771
9.016 | -34.4303 | -15.7808
9.0221 | -34.4307 | -15.7905
9.0282 | -34.4311 | -15.8002
9.0343 | -34.4315 | -15.8101
9.0404 | -34.432 | -15.8199
9.0465 | -34.4323 | -15.8298
9.0526 | -34.4328 | -15.8395
9.0587 | -34.4332 | -15.8494
9.0648 | -34.4336 | -15.859
9.0709 | -34.434 | -15.8688
9.077 | -34.4344 | -15.8786
9.0831 | -34.4348 | -15.8885
9.0892 | -34.4352 | -15.8984
9.0953 | -34.4357 | -15.9083
9.1015 | -34.4361 | -15.9181
9.1076 | -34.4365 | -15.9276
9.1137 | -34.4369 | -15.9374
9.1198 | -34.4373 | -15.9472
9.1259 | -34.4377 | -15.957
9.132 | -34.4381 | -15.9669
9.1381 | -34.4385 | -15.9765
9.1442 | -34.439 | -15.9859
9.1503 | -34.4394 | -15.9957
9.1564 | -34.4398 | -16.0055
9.1625 | -34.4402 | -16.0152
9.1686 | -34.4406 | -16.0249
9.1747 | -34.441 | -16.0347
9.1808 | -34.4414 | -16.0444
9.1869 | -34.4418 | -16.0542
9.193 | -34.4423 | -16.064
9.1991 | -34.4426 | -16.0738
9.2052 | -34.4431 | -16.0835
9.2113 | -34.4435 | -16.0933
9.2174 | -34.444 | -16.1029
9.2235 | -34.4444 | -16.1126
9.2296 | -34.4448 | -16.1223
9.2357 | -34.4452 | -16.1322
9.2419 | -34.4456 | -16.1419
9.248 | -34.4461 | -16.1517
9.2541 | -34.4465 | -16.1615
9.2602 | -34.4469 | -16.1713
9.2663 | -34.4473 | -16.1809
9.2724 | -34.4478 | -16.1905
9.2785 | -34.4482 | -16.2002
9.2846 | -34.4486 | -16.2098
9.2907 | -34.449 | -16.2195
9.2968 | -34.4494 | -16.2294
9.3029 | -34.4498 | -16.2393
9.309 | -34.4503 | -16.249
9.3151 | -34.4507 | -16.2586
9.3212 | -34.4511 | -16.2682
9.3273 | -34.4515 | -16.2779
9.3334 | -34.4519 | -16.2875
9.3395 | -34.4523 | -16.2971
9.3456 | -34.4527 | -16.3066
9.3517 | -34.4531 | -16.3164
9.3578 | -34.4536 | -16.326
9.3639 | -34.454 | -16.3356
9.37 | -34.4544 | -16.3454
9.3761 | -34.4549 | -16.3552
9.3822 | -34.4553 | -16.3649
9.3884 | -34.4557 | -16.3742
9.3945 | -34.4561 | -16.384
9.4006 | -34.4566 | -16.3938
9.4067 | -34.457 | -16.4035
9.4128 | -34.4574 | -16.4133
9.4189 | -34.4578 | -16.4231
9.425 | -34.4583 | -16.4327
9.4311 | -34.4587 | -16.4425
9.4372 | -34.4592 | -16.4521
9.4433 | -34.4596 | -16.4618
9.4494 | -34.46 | -16.4714
9.4555 | -34.4605 | -16.4811
9.4616 | -34.4609 | -16.4908
9.4677 | -34.4614 | -16.5005
9.4738 | -34.4618 | -16.5101
9.4799 | -34.4622 | -16.5198
9.486 | -34.4626 | -16.5295
9.4921 | -34.463 | -16.5391
9.4982 | -34.4635 | -16.5486
9.5043 | -34.4639 | -16.5583
9.5104 | -34.4644 | -16.5678
9.5165 | -34.4648 | -16.5775
9.5226 | -34.4652 | -16.5874
9.5288 | -34.4656 | -16.5969
9.5349 | -34.466 | -16.6066
9.541 | -34.4665 | -16.616
9.5471 | -34.4669 | -16.6256
9.5532 | -34.4673 | -16.6353
9.5593 | -34.4678 | -16.6449
9.5654 | -34.4682 | -16.6545
9.5715 | -34.4686 | -16.6641
9.5776 | -34.469 | -16.674
9.5837 | -34.4694 | -16.6836
9.5898 | -34.4698 | -16.6933
9.5959 | -34.4703 | -16.703
9.602 | -34.4707 | -16.7126
9.6081 | -34.4712 | -16.7222
9.6142 | -34.4716 | -16.7318
9.6203 | -34.472 | -16.7415
9.6264 | -34.4724 | -16.751
9.6325 | -34.4728 | -16.7607
9.6386 | -34.4733 | -16.7702
9.6447 | -34.4737 | -16.7799
9.6508 | -34.4741 | -16.7894
9.6569 | -34.4745 | -16.7992
9.663 | -34.475 | -16.8088
9.6691 | -34.4754 | -16.8185
9.6753 | -34.4759 | -16.8281
9.6814 | -34.4763 | -16.8375
9.6875 | -34.4768 | -16.847
9.6936 | -34.4772 | -16.8564
9.6997 | -34.4777 | -16.8661
9.7058 | -34.4781 | -16.8756
9.7119 | -34.4786 | -16.8853
9.718 | -34.479 | -16.8951
9.7241 | -34.4794 | -16.9046
9.7302 | -34.4798 | -16.9143
9.7363 | -34.4803 | -16.9238
9.7424 | -34.4808 | -16.9333
9.7485 | -34.4812 | -16.9429
9.7546 | -34.4816 | -16.9525
9.7607 | -34.482 | -16.9623
9.7668 | -34.4825 | -16.972
9.7729 | -34.4829 | -16.9815
9.779 | -34.4834 | -16.9908
9.7851 | -34.4839 | -17.0006
9.7912 | -34.4843 | -17.0102
9.7973 | -34.4848 | -17.02
9.8034 | -34.4852 | -17.0295
9.8095 | -34.4857 | -17.0392
9.8157 | -34.4861 | -17.0489
9.8218 | -34.4866 | -17.0587
9.8279 | -34.487 | -17.0681
9.834 | -34.4874 | -17.0778
9.8401 | -34.4879 | -17.0874
9.8462 | -34.4883 | -17.0968
9.8523 | -34.4887 | -17.1065
9.8584 | -34.4892 | -17.116
9.8645 | -34.4896 | -17.1257
9.8706 | -34.49 | -17.1353
9.8767 | -34.4904 | -17.1449
9.8828 | -34.4908 | -17.1544
9.8889 | -34.4913 | -17.1639
9.895 | -34.4917 | -17.1735
9.9011 | -34.4922 | -17.183
9.9072 | -34.4926 | -17.1925
9.9133 | -34.493 | -17.2022
9.9194 | -34.4935 | -17.2117
9.9255 | -34.4939 | -17.2212
9.9316 | -34.4944 | -17.2307
9.9377 | -34.4948 | -17.2403
9.9438 | -34.4952 | -17.2499
9.9499 | -34.4957 | -17.2597
9.956 | -34.4961 | -17.2692
9.9622 | -34.4966 | -17.2788
9.9683 | -34.497 | -17.2883
9.9744 | -34.4974 | -17.2981
9.9805 | -34.4979 | -17.3076
9.9866 | -34.4983 | -17.317
9.9927 | -34.4987 | -17.3265
9.9988 | -34.4991 | -17.336
10.0049 | -34.4996 | -17.3455
10.011 | -34.5 | -17.3552
10.0171 | -34.5005 | -17.3645
10.0232 | -34.5009 | -17.374
10.0293 | -34.5014 | -17.3837
10.0354 | -34.5018 | -17.3933
10.0415 | -34.5023 | -17.4029
10.0476 | -34.5027 | -17.4126
10.0537 | -34.5032 | -17.4223
10.0598 | -34.5036 | -17.4319
10.0659 | -34.504 | -17.4416
10.072 | -34.5045 | -17.4509
10.0781 | -34.5049 | -17.4604
10.0842 | -34.5054 | -17.47
10.0903 | -34.5058 | -17.4797
10.0964 | -34.5063 | -17.4891
10.1026 | -34.5067 | -17.4988
10.1087 | -34.5071 | -17.5084
10.1148 | -34.5076 | -17.5181
10.1209 | -34.508 | -17.5274
10.127 | -34.5084 | -17.5366
10.1331 | -34.5089 | -17.5462
10.1392 | -34.5093 | -17.5556
10.1453 | -34.5097 | -17.565
10.1514 | -34.5102 | -17.5747
10.1575 | -34.5106 | -17.5845
10.1636 | -34.511 | -17.5938
10.1697 | -34.5115 | -17.6034
10.1758 | -34.5119 | -17.6128
10.1819 | -34.5123 | -17.6223
10.188 | -34.5128 | -17.6318
10.1941 | -34.5132 | -17.6413
10.2002 | -34.5136 | -17.6508
10.2063 | -34.5141 | -17.6603
10.2124 | -34.5145 | -17.6699
10.2185 | -34.515 | -17.6795
10.2246 | -34.5154 | -17.6889
10.2307 | -34.5158 | -17.6985
10.2368 | -34.5163 | -17.7077
10.2429 | -34.5167 | -17.7173
10.2491 | -34.5171 | -17.7269
10.2552 | -34.5176 | -17.7365
10.2613 | -34.518 | -17.7461
10.2674 | -34.5185 | -17.7556
10.2735 | -34.5189 | -17.7653
10.2796 | -34.5194 | -17.7749
10.2857 | -34.5198 | -17.7844
10.2918 | -34.5202 | -17.7939
10.2979 | -34.5207 | -17.8033
10.304 | -34.5211 | -17.8127
10.3101 | -34.5216 | -17.8224
10.3162 | -34.522 | -17.8321
10.3223 | -34.5225 | -17.8413
10.3284 | -34.5229 | -17.8508
10.3345 | -34.5234 | -17.8602
10.3406 | -34.5238 | -17.8697
10.3467 | -34.5242 | -17.8793
10.3528 | -34.5246 | -17.8888
10.3589 | -34.5251 | -17.8983
10.365 | -34.5255 | -17.908
10.3711 | -34.526 | -17.9177
10.3772 | -34.5264 | -17.9272
10.3833 | -34.5269 | -17.9367
10.3895 | -34.5273 | -17.946
10.3956 | -34.5277 | -17.9554
10.4017 | -34.5282 | -17.9646
10.4078 | -34.5286 | -17.974
10.4139 | -34.529 | -17.9836
10.42 | -34.5294 | -17.9934
10.4261 | -34.5299 | -18.0029
10.4322 | -34.5303 | -18.0125
10.4383 | -34.5307 | -18.022
10.4444 | -34.5312 | -18.0313
10.4505 | -34.5316 | -18.0409
10.4566 | -34.532 | -18.0502
10.4627 | -34.5324 | -18.0595
10.4688 | -34.5329 | -18.0689
10.4749 | -34.5333 | -18.0785
10.481 | -34.5337 | -18.0879
10.4871 | -34.5342 | -18.0973
10.4932 | -34.5346 | -18.1066
10.4993 | -34.5351 | -18.1161
10.5054 | -34.5355 | -18.1256
10.5115 | -34.5359 | -18.1352
10.5176 | -34.5363 | -18.1447
10.5237 | -34.5368 | -18.1543
10.5298 | -34.5372 | -18.1636
10.536 | -34.5376 | -18.1732
10.5421 | -34.5381 | -18.1826
10.5482 | -34.5385 | -18.1921
10.5543 | -34.539 | -18.2014
10.5604 | -34.5394 | -18.2106
10.5665 | -34.5398 | -18.2199
10.5726 | -34.5403 | -18.2294
10.5787 | -34.5407 | -18.2387
10.5848 | -34.5411 | -18.2482
10.5909 | -34.5416 | -18.2576
10.597 | -34.542 | -18.2672
10.6031 | -34.5425 | -18.2768
10.6092 | -34.5429 | -18.2862
10.6153 | -34.5433 | -18.2958
10.6214 | -34.5438 | -18.3053
10.6275 | -34.5442 | -18.3147
10.6336 | -34.5447 | -18.3242
10.6397 | -34.5451 | -18.3336
10.6458 | -34.5455 | -18.3429
10.6519 | -34.5459 | -18.3525
10.658 | -34.5464 | -18.3621
10.6641 | -34.5468 | -18.3718
10.6702 | -34.5472 | -18.3813
10.6764 | -34.5477 | -18.3908
10.6825 | -34.5481 | -18.4001
10.6886 | -34.5485 | -18.4094
10.6947 | -34.549 | -18.4185
10.7008 | -34.5494 | -18.4279
10.7069 | -34.5498 | -18.4374
10.713 | -34.5503 | -18.4468
10.7191 | -34.5507 | -18.4563
10.7252 | -34.5511 | -18.4657
10.7313 | -34.5516 | -18.4751
10.7374 | -34.552 | -18.4845
10.7435 | -34.5524 | -18.4939
10.7496 | -34.5529 | -18.503
10.7557 | -34.5533 | -18.5125
10.7618 | -34.5537 | -18.5218
10.7679 | -34.5541 | -18.5311
10.774 | -34.5545 | -18.5406
10.7801 | -34.555 | -18.5501
10.7862 | -34.5554 | -18.5593
10.7923 | -34.5558 | -18.5688
10.7984 | -34.5563 | -18.578
10.8045 | -34.5568 | -18.5871
10.8106 | -34.5572 | -18.5967
10.8168 | -34.5576 | -18.6061
10.8229 | -34.558 | -18.6155
10.829 | -34.5584 | -18.625
10.8351 | -34.5589 | -18.6341
10.8412 | -34.5593 | -18.6434
10.8473 | -34.5598 | -18.6526
10.8534 | -34.5603 | -18.6621
10.8595 | -34.5607 | -18.6714
10.8656 | -34.5611 | -18.6807
10.8717 | -34.5616 | -18.6901
10.8778 | -34.562 | -18.6993
10.8839 | -34.5624 | -18.7088
10.89 | -34.5629 | -18.7181
10.8961 | -34.5634 | -18.7273
10.9022 | -34.5638 | -18.7365
10.9083 | -34.5642 | -18.7458
10.9144 | -34.5646 | -18.755
10.9205 | -34.565 | -18.7644
10.9266 | -34.5655 | -18.7737
10.9327 | -34.5659 | -18.783
10.9388 | -34.5663 | -18.7919
10.9449 | -34.5668 | -18.8011
10.951 | -34.5672 | -18.8104
10.9571 | -34.5677 | -18.8197
10.9633 | -34.5681 | -18.8291
10.9694 | -34.5685 | -18.8386
10.9755 | -34.569 | -18.8478
10.9816 | -34.5694 | -18.8572
10.9877 | -34.5699 | -18.8665
10.9938 | -34.5703 | -18.8758
10.9999 | -34.5708 | -18.8851
11.006 | -34.5712 | -18.8942
11.0121 | -34.5716 | -18.9035
11.0182 | -34.572 | -18.9128
11.0243 | -34.5725 | -18.922
11.0304 | -34.5729 | -18.9314
11.0365 | -34.5733 | -18.9405
11.0426 | -34.5738 | -18.9499
11.0487 | -34.5742 | -18.9595
11.0548 | -34.5747 | -18.9687
11.0609 | -34.5752 | -18.9779
11.067 | -34.5756 | -18.9873
11.0731 | -34.5761 | -18.9964
11.0792 | -34.5765 | -19.0056
11.0853 | -34.577 | -19.0149
11.0914 | -34.5774 | -19.0243
11.0975 | -34.5779 | -19.0337
11.1037 | -34.5783 | -19.0429
11.1098 | -34.5787 | -19.0523
11.1159 | -34.5791 | -19.0613
11.122 | -34.5795 | -19.0704
11.1281 | -34.58 | -19.0796
11.1342 | -34.5804 | -19.0889
11.1403 | -34.5809 | -19.0982
11.1464 | -34.5813 | -19.1077
11.1525 | -34.5818 | -19.1165
11.1586 | -34.5822 | -19.1258
11.1647 | -34.5827 | -19.1351
11.1708 | -34.5831 | -19.1444
11.1769 | -34.5835 | -19.1538
11.183 | -34.584 | -19.1629
11.1891 | -34.5845 | -19.1721
11.1952 | -34.5849 | -19.1814
11.2013 | -34.5853 | -19.1908
11.2074 | -34.5858 | -19.2001
11.2135 | -34.5862 | -19.2096
11.2196 | -34.5867 | -19.2185
11.2257 | -34.5872 | -19.2279
11.2318 | -34.5876 | -19.2372
11.2379 | -34.5881 | -19.2464
11.244 | -34.5885 | -19.2557
11.2502 | -34.589 | -19.2649
11.2563 | -34.5895 | -19.2742
11.2624 | -34.5899 | -19.2833
11.2685 | -34.5904 | -19.2925
11.2746 | -34.5908 | -19.3016
11.2807 | -34.5912 | -19.3109
11.2868 | -34.5917 | -19.3204
11.2929 | -34.5922 | -19.3292
11.299 | -34.5926 | -19.3385
11.3051 | -34.5931 | -19.3478
11.3112 | -34.5935 | -19.3569
11.3173 | -34.594 | -19.3661
11.3234 | -34.5945 | -19.3748
11.3295 | -34.5949 | -19.3841
11.3356 | -34.5953 | -19.3933
11.3417 | -34.5957 | -19.4024
11.3478 | -34.5962 | -19.4117
11.3539 | -34.5966 | -19.4209
11.36 | -34.597 | -19.4297
11.3661 | -34.5975 | -19.4387
11.3722 | -34.598 | -19.4477
11.3783 | -34.5984 | -19.457
11.3844 | -34.5989 | -19.4663
11.3906 | -34.5993 | -19.4755
11.3967 | -34.5998 | -19.4845
11.4028 | -34.6002 | -19.4938
11.4089 | -34.6007 | -19.5029
11.415 | -34.6011 | -19.512
11.4211 | -34.6016 | -19.5213
11.4272 | -34.602 | -19.5306
11.4333 | -34.6025 | -19.5393
11.4394 | -34.6029 | -19.5483
11.4455 | -34.6033 | -19.5575
11.4516 | -34.6037 | -19.5666
11.4577 | -34.6042 | -19.5758
11.4638 | -34.6047 | -19.5846
11.4699 | -34.6051 | -19.5938
11.476 | -34.6056 | -19.6031
11.4821 | -34.606 | -19.6122
11.4882 | -34.6065 | -19.6214
11.4943 | -34.6069 | -19.6306
11.5004 | -34.6074 | -19.6395
11.5065 | -34.6078 | -19.6484
11.5126 | -34.6083 | -19.6574
11.5187 | -34.6087 | -19.6667
11.5248 | -34.6092 | -19.676
11.5309 | -34.6096 | -19.6851
11.5371 | -34.6101 | -19.6938
11.5432 | -34.6106 | -19.703
11.5493 | -34.611 | -19.7122
11.5554 | -34.6115 | -19.7214
11.5615 | -34.6119 | -19.7305
11.5676 | -34.6124 | -19.7396
11.5737 | -34.6129 | -19.7484
11.5798 | -34.6134 | -19.7575
11.5859 | -34.6139 | -19.7666
11.592 | -34.6143 | -19.7756
11.5981 | -34.6147 | -19.7847
11.6042 | -34.6152 | -19.7938
11.6103 | -34.6157 | -19.8028
11.6164 | -34.6161 | -19.8118
11.6225 | -34.6166 | -19.8208
11.6286 | -34.617 | -19.8299
11.6347 | -34.6175 | -19.8391
11.6408 | -34.6179 | -19.8479
11.6469 | -34.6184 | -19.857
11.653 | -34.6188 | -19.866
11.6591 | -34.6193 | -19.875
11.6652 | -34.6197 | -19.8839
11.6713 | -34.6202 | -19.8931
11.6775 | -34.6207 | -19.9018
11.6836 | -34.6211 | -19.9108
11.6897 | -34.6216 | -19.9199
11.6958 | -34.6221 | -19.9291
11.7019 | -34.6225 | -19.9383
11.708 | -34.623 | -19.9472
11.7141 | -34.6235 | -19.956
11.7202 | -34.6239 | -19.9651
11.7263 | -34.6243 | -19.9742
11.7324 | -34.6248 | -19.9833
11.7385 | -34.6252 | -19.9923
11.7446 | -34.6257 | -20.0013
11.7507 | -34.6262 | -20.01
11.7568 | -34.6266 | -20.019
11.7629 | -34.6271 | -20.0282
11.769 | -34.6275 | -20.0371
11.7751 | -34.628 | -20.0461
11.7812 | -34.6285 | -20.0549
11.7873 | -34.6289 | -20.064
11.7934 | -34.6294 | -20.0729
11.7995 | -34.6299 | -20.0818
11.8056 | -34.6303 | -20.0906
11.8117 | -34.6308 | -20.0996
11.8178 | -34.6313 | -20.1083
11.824 | -34.6317 | -20.1172
11.8301 | -34.6322 | -20.1261
11.8362 | -34.6326 | -20.135
11.8423 | -34.6331 | -20.1441
11.8484 | -34.6335 | -20.153
11.8545 | -34.634 | -20.1617
11.8606 | -34.6345 | -20.1705
11.8667 | -34.6349 | -20.1796
11.8728 | -34.6354 | -20.1884
11.8789 | -34.6359 | -20.1973
11.885 | -34.6363 | -20.2063
11.8911 | -34.6368 | -20.2152
11.8972 | -34.6373 | -20.2239
11.9033 | -34.6377 | -20.2327
11.9094 | -34.6382 | -20.2415
11.9155 | -34.6387 | -20.2504
11.9216 | -34.6392 | -20.259
11.9277 | -34.6396 | -20.2678
11.9338 | -34.6401 | -20.2768
11.9399 | -34.6405 | -20.2858
11.946 | -34.641 | -20.2947
11.9521 | -34.6415 | -20.3035
11.9582 | -34.642 | -20.3123
11.9644 | -34.6424 | -20.321
11.9705 | -34.6429 | -20.3299
11.9766 | -34.6434 | -20.3388
11.9827 | -34.6439 | -20.3476
11.9888 | -34.6444 | -20.3563
11.9949 | -34.6449 | -20.3648
12.001 | -34.6454 | -20.3738
12.0071 | -34.6459 | -20.3826
12.0132 | -34.6463 | -20.3913
12.0193 | -34.6468 | -20.4002
12.0254 | -34.6472 | -20.409
12.0315 | -34.6477 | -20.4175
12.0376 | -34.6482 | -20.4263
12.0437 | -34.6486 | -20.4351
12.0498 | -34.6491 | -20.444
12.0559 | -34.6496 | -20.4527
12.062 | -34.6501 | -20.4613
12.0681 | -34.6506 | -20.4702
12.0742 | -34.651 | -20.4789
12.0803 | -34.6515 | -20.4879
12.0864 | -34.652 | -20.4968
12.0925 | -34.6524 | -20.5056
12.0986 | -34.653 | -20.5144
12.1047 | -34.6534 | -20.5232
12.1109 | -34.6539 | -20.5319
12.117 | -34.6544 | -20.5406
12.1231 | -34.6549 | -20.5494
12.1292 | -34.6553 | -20.5581
12.1353 | -34.6558 | -20.5669
12.1414 | -34.6563 | -20.5757
12.1475 | -34.6568 | -20.5846
12.1536 | -34.6572 | -20.5934
12.1597 | -34.6577 | -20.6022
12.1658 | -34.6582 | -20.6109
12.1719 | -34.6587 | -20.6195
12.178 | -34.6591 | -20.6283
12.1841 | -34.6596 | -20.6373
12.1902 | -34.6601 | -20.646
12.1963 | -34.6606 | -20.6548
12.2024 | -34.6611 | -20.6634
12.2085 | -34.6616 | -20.6722
12.2146 | -34.6621 | -20.681
12.2207 | -34.6626 | -20.6897
12.2268 | -34.663 | -20.6985
12.2329 | -34.6635 | -20.7071
12.239 | -34.664 | -20.7155
12.2451 | -34.6645 | -20.7242
12.2513 | -34.665 | -20.7331
12.2574 | -34.6655 | -20.7418
12.2635 | -34.666 | -20.7506
12.2696 | -34.6664 | -20.7593
12.2757 | -34.6669 | -20.7678
12.2818 | -34.6674 | -20.7766
12.2879 | -34.6679 | -20.7854
12.294 | -34.6684 | -20.7939
12.3001 | -34.6689 | -20.8027
12.3062 | -34.6694 | -20.8114
12.3123 | -34.6699 | -20.8198
12.3184 | -34.6704 | -20.8284
12.3245 | -34.6708 | -20.8369
12.3306 | -34.6713 | -20.8456
12.3367 | -34.6718 | -20.8543
12.3428 | -34.6723 | -20.863
12.3489 | -34.6728 | -20.8715
12.355 | -34.6733 | -20.8803
12.3611 | -34.6738 | -20.8891
12.3672 | -34.6742 | -20.8977
12.3733 | -34.6747 | -20.9064
12.3794 | -34.6752 | -20.9147
12.3855 | -34.6757 | -20.9233
12.3916 | -34.6762 | -20.9319
12.3978 | -34.6767 | -20.9408
12.4039 | -34.6772 | -20.9495
12.41 | -34.6777 | -20.9581
12.4161 | -34.6782 | -20.9665
12.4222 | -34.6787 | -20.9751
12.4283 | -34.6791 | -20.9837
12.4344 | -34.6796 | -20.9924
12.4405 | -34.6801 | -21.0012
12.4466 | -34.6806 | -21.0097
12.4527 | -34.6811 | -21.0182
12.4588 | -34.6816 | -21.0269
12.4649 | -34.682 | -21.0355
12.471 | -34.6826 | -21.0444
12.4771 | -34.683 | -21.053
12.4832 | -34.6835 | -21.0616
12.4893 | -34.684 | -21.0701
12.4954 | -34.6845 | -21.0786
12.5015 | -34.685 | -21.0872
12.5076 | -34.6855 | -21.096
12.5137 | -34.686 | -21.1046
12.5198 | -34.6865 | -21.1131
12.5259 | -34.687 | -21.1218
12.532 | -34.6874 | -21.1304
12.5382 | -34.688 | -21.1391
12.5443 | -34.6884 | -21.1477
12.5504 | -34.6889 | -21.1564
12.5565 | -34.6894 | -21.1649
12.5626 | -34.6899 | -21.1737
12.5687 | -34.6904 | -21.1823
12.5748 | -34.6909 | -21.1909
12.5809 | -34.6914 | -21.1995
12.587 | -34.6919 | -21.2082
12.5931 | -34.6924 | -21.2167
12.5992 | -34.693 | -21.2253
12.6053 | -34.6934 | -21.2339
12.6114 | -34.6939 | -21.2425
12.6175 | -34.6944 | -21.2512
12.6236 | -34.6949 | -21.2599
12.6297 | -34.6954 | -21.2685
12.6358 | -34.6959 | -21.2769
12.6419 | -34.6964 | -21.2855
12.648 | -34.6969 | -21.2941
12.6541 | -34.6974 | -21.3027
12.6602 | -34.6979 | -21.3109
12.6663 | -34.6984 | -21.3196
12.6724 | -34.6989 | -21.328
12.6785 | -34.6994 | -21.3366
12.6847 | -34.6999 | -21.3451
12.6908 | -34.7004 | -21.3537
12.6969 | -34.7009 | -21.3622
12.703 | -34.7014 | -21.3709
12.7091 | -34.7019 | -21.3795
12.7152 | -34.7024 | -21.3879
12.7213 | -34.7029 | -21.3966
12.7274 | -34.7034 | -21.4052
12.7335 | -34.7039 | -21.4135
12.7396 | -34.7044 | -21.4219
12.7457 | -34.7049 | -21.4305
12.7518 | -34.7053 | -21.4391
12.7579 | -34.7058 | -21.4478
12.764 | -34.7063 | -21.4564
12.7701 | -34.7068 | -21.4651
12.7762 | -34.7073 | -21.4736
12.7823 | -34.7078 | -21.4822
12.7884 | -34.7083 | -21.4907
12.7945 | -34.7088 | -21.4995
12.8006 | -34.7093 | -21.5081
12.8067 | -34.7098 | -21.5165
12.8128 | -34.7103 | -21.5249
12.8189 | -34.7108 | -21.5334
12.8251 | -34.7112 | -21.5419
12.8312 | -34.7117 | -21.5505
12.8373 | -34.7122 | -21.5591
12.8434 | -34.7127 | -21.5677
12.8495 | -34.7132 | -21.5762
12.8556 | -34.7137 | -21.5847
12.8617 | -34.7142 | -21.5933
12.8678 | -34.7147 | -21.6021
12.8739 | -34.7152 | -21.6107
12.88 | -34.7156 | -21.6192
12.8861 | -34.7161 | -21.6276
12.8922 | -34.7166 | -21.6362
12.8983 | -34.7171 | -21.6447
12.9044 | -34.7176 | -21.6531
12.9105 | -34.718 | -21.6617
12.9166 | -34.7185 | -21.6702
12.9227 | -34.719 | -21.6788
12.9288 | -34.7195 | -21.6874
12.9349 | -34.72 | -21.6959
12.941 | -34.7205 | -21.7045
12.9471 | -34.721 | -21.7131
12.9532 | -34.7215 | -21.7216
12.9593 | -34.722 | -21.7304
12.9654 | -34.7225 | -21.7389
12.9716 | -34.723 | -21.7477
12.9777 | -34.7235 | -21.7563
12.9838 | -34.724 | -21.7649
12.9899 | -34.7245 | -21.7735
12.996 | -34.7249 | -21.7821
13.0021 | -34.7255 | -21.7906
13.0082 | -34.726 | -21.799
13.0143 | -34.7265 | -21.8074
13.0204 | -34.727 | -21.8158
13.0265 | -34.7275 | -21.8243
13.0326 | -34.7279 | -21.8329
13.0387 | -34.7285 | -21.8415
13.0448 | -34.7289 | -21.8501
13.0509 | -34.7294 | -21.8586
13.057 | -34.7299 | -21.867
13.0631 | -34.7305 | -21.8755
13.0692 | -34.731 | -21.8839
13.0753 | -34.7314 | -21.8926
13.0814 | -34.7319 | -21.9013
13.0875 | -34.7324 | -21.9096
13.0936 | -34.7329 | -21.9183
13.0997 | -34.7334 | -21.9267
13.1058 | -34.7339 | -21.9351
13.112 | -34.7344 | -21.9436
13.1181 | -34.7349 | -21.9521
13.1242 | -34.7354 | -21.9607
13.1303 | -34.7359 | -21.9691
13.1364 | -34.7363 | -21.9776
13.1425 | -34.7368 | -21.986
13.1486 | -34.7373 | -21.9945
13.1547 | -34.7378 | -22.003
13.1608 | -34.7383 | -22.0115
13.1669 | -34.7388 | -22.0199
13.173 | -34.7393 | -22.0285
13.1791 | -34.7398 | -22.0369
13.1852 | -34.7403 | -22.0454
13.1913 | -34.7407 | -22.0538
13.1974 | -34.7412 | -22.0625
13.2035 | -34.7417 | -22.0709
13.2096 | -34.7422 | -22.0793
13.2157 | -34.7427 | -22.0876
13.2218 | -34.7432 | -22.0961
13.2279 | -34.7437 | -22.1045
13.234 | -34.7442 | -22.113
13.2401 | -34.7447 | -22.1215
13.2462 | -34.7452 | -22.13
13.2524 | -34.7457 | -22.1385
13.2585 | -34.7461 | -22.147
13.2646 | -34.7466 | -22.1555
13.2707 | -34.7472 | -22.1641
13.2768 | -34.7477 | -22.1727
13.2829 | -34.7482 | -22.1814
13.289 | -34.7487 | -22.1897
13.2951 | -34.7492 | -22.198
13.3012 | -34.7497 | -22.2066
13.3073 | -34.7502 | -22.2153
13.3134 | -34.7506 | -22.2237
13.3195 | -34.7511 | -22.2322
13.3256 | -34.7516 | -22.2406
13.3317 | -34.752 | -22.2491
13.3378 | -34.7525 | -22.2576
13.3439 | -34.753 | -22.2663
13.35 | -34.7535 | -22.2749
13.3561 | -34.754 | -22.2833
13.3622 | -34.7545 | -22.2917
13.3683 | -34.755 | -22.3001
13.3744 | -34.7555 | -22.3085
13.3805 | -34.756 | -22.3171
13.3866 | -34.7564 | -22.3255
13.3927 | -34.7569 | -22.3342
13.3989 | -34.7574 | -22.3426
13.405 | -34.7579 | -22.351
13.4111 | -34.7584 | -22.3594
13.4172 | -34.7588 | -22.3679
13.4233 | -34.7593 | -22.3766
13.4294 | -34.7598 | -22.3852
13.4355 | -34.7603 | -22.3935
13.4416 | -34.7608 | -22.402
13.4477 | -34.7613 | -22.4104
13.4538 | -34.7618 | -22.4189
13.4599 | -34.7622 | -22.4275
13.466 | -34.7627 | -22.4361
13.4721 | -34.7632 | -22.4444
13.4782 | -34.7637 | -22.4528
13.4843 | -34.7642 | -22.4615
13.4904 | -34.7647 | -22.4698
13.4965 | -34.7651 | -22.4784
13.5026 | -34.7656 | -22.4869
13.5087 | -34.7661 | -22.4953
13.5148 | -34.7666 | -22.5038
13.5209 | -34.7671 | -22.5125
13.527 | -34.7675 | -22.521
13.5331 | -34.768 | -22.5293
13.5393 | -34.7685 | -22.5377
13.5454 | -34.769 | -22.546
13.5515 | -34.7695 | -22.5547
13.5576 | -34.77 | -22.5633
13.5637 | -34.7705 | -22.5719
13.5698 | -34.771 | -22.5804
13.5759 | -34.7715 | -22.5886
13.582 | -34.7719 | -22.597
13.5881 | -34.7724 | -22.6056
13.5942 | -34.7729 | -22.6142
13.6003 | -34.7734 | -22.6226
13.6064 | -34.7739 | -22.6313
13.6125 | -34.7744 | -22.6399
13.6186 | -34.7749 | -22.6482
13.6247 | -34.7754 | -22.6568
13.6308 | -34.7759 | -22.6654
13.6369 | -34.7764 | -22.6739
13.643 | -34.7768 | -22.6825
13.6491 | -34.7773 | -22.6908
13.6552 | -34.7778 | -22.6993
13.6613 | -34.7783 | -22.7077
13.6674 | -34.7788 | -22.7163
13.6735 | -34.7793 | -22.7248
13.6796 | -34.7798 | -22.7332
13.6858 | -34.7802 | -22.7415
13.6919 | -34.7807 | -22.75
13.698 | -34.7812 | -22.7585
13.7041 | -34.7817 | -22.7671
13.7102 | -34.7822 | -22.7756
13.7163 | -34.7827 | -22.784
13.7224 | -34.7832 | -22.7924
13.7285 | -34.7837 | -22.8007
13.7346 | -34.7841 | -22.8093
13.7407 | -34.7846 | -22.8178
13.7468 | -34.7851 | -22.8262
13.7529 | -34.7856 | -22.8347
13.759 | -34.786 | -22.8432
13.7651 | -34.7865 | -22.8517
13.7712 | -34.787 | -22.8602
13.7773 | -34.7875 | -22.8687
13.7834 | -34.7879 | -22.8772
13.7895 | -34.7884 | -22.8855
13.7956 | -34.7889 | -22.8941
13.8017 | -34.7894 | -22.9025
13.8078 | -34.7899 | -22.9109
13.8139 | -34.7903 | -22.9193
13.82 | -34.7908 | -22.9276
13.8262 | -34.7913 | -22.936
13.8323 | -34.7918 | -22.9446
13.8384 | -34.7923 | -22.9532
13.8445 | -34.7928 | -22.9617
13.8506 | -34.7932 | -22.9701
13.8567 | -34.7937 | -22.9787
13.8628 | -34.7942 | -22.9872
13.8689 | -34.7947 | -22.9956
13.875 | -34.7952 | -23.0041
13.8811 | -34.7956 | -23.0126
13.8872 | -34.7961 | -23.021
13.8933 | -34.7966 | -23.0293
13.8994 | -34.797 | -23.0378
13.9055 | -34.7975 | -23.0465
13.9116 | -34.798 | -23.0548
13.9177 | -34.7985 | -23.0632
13.9238 | -34.7989 | -23.0717
13.9299 | -34.7994 | -23.08
13.936 | -34.7999 | -23.0885
13.9421 | -34.8004 | -23.0968
13.9482 | -34.8008 | -23.1051
13.9543 | -34.8013 | -23.1136
13.9604 | -34.8018 | -23.122
13.9665 | -34.8023 | -23.1305
13.9727 | -34.8028 | -23.1392
13.9788 | -34.8033 | -23.1476
13.9849 | -34.8038 | -23.1561
13.991 | -34.8042 | -23.1646
13.9971 | -34.8047 | -23.1728
14.0032 | -34.8051 | -23.1812
14.0093 | -34.8056 | -23.1896
14.0154 | -34.8061 | -23.198
14.0215 | -34.8066 | -23.2064
14.0276 | -34.8071 | -23.2148
14.0337 | -34.8076 | -23.2232
14.0398 | -34.808 | -23.2316
14.0459 | -34.8085 | -23.24
14.052 | -34.809 | -23.2485
14.0581 | -34.8095 | -23.257
14.0642 | -34.81 | -23.2654
14.0703 | -34.8104 | -23.2738
14.0764 | -34.8109 | -23.2822
14.0825 | -34.8114 | -23.2906
14.0886 | -34.8119 | -23.299
14.0947 | -34.8124 | -23.3075
14.1008 | -34.8128 | -23.3159
14.1069 | -34.8133 | -23.324
14.1131 | -34.8138 | -23.3326
14.1192 | -34.8143 | -23.3409
14.1253 | -34.8147 | -23.3491
14.1314 | -34.8152 | -23.3576
14.1375 | -34.8157 | -23.366
14.1436 | -34.8162 | -23.3743
14.1497 | -34.8167 | -23.3827
14.1558 | -34.8172 | -23.3911
14.1619 | -34.8176 | -23.3995
14.168 | -34.8181 | -23.4079
14.1741 | -34.8186 | -23.4162
14.1802 | -34.8191 | -23.4245
14.1863 | -34.8196 | -23.4329
14.1924 | -34.8201 | -23.4413
14.1985 | -34.8206 | -23.4498
14.2046 | -34.8211 | -23.4582
14.2107 | -34.8216 | -23.4664
14.2168 | -34.8221 | -23.4749
14.2229 | -34.8225 | -23.4834
14.229 | -34.823 | -23.4918
14.2351 | -34.8235 | -23.5003
14.2412 | -34.824 | -23.5087
14.2473 | -34.8245 | -23.5171
14.2534 | -34.825 | -23.5253
14.2596 | -34.8254 | -23.5335
14.2657 | -34.8259 | -23.542
14.2718 | -34.8264 | -23.5503
14.2779 | -34.8269 | -23.5587
14.284 | -34.8274 | -23.5671
14.2901 | -34.8279 | -23.5754
14.2962 | -34.8283 | -23.5838
14.3023 | -34.8288 | -23.5922
14.3084 | -34.8293 | -23.6006
14.3145 | -34.8298 | -23.609
14.3206 | -34.8303 | -23.6175
14.3267 | -34.8308 | -23.6258
14.3328 | -34.8313 | -23.6342
14.3389 | -34.8317 | -23.6425
14.345 | -34.8322 | -23.6507
14.3511 | -34.8327 | -23.6591
14.3572 | -34.8332 | -23.6676
14.3633 | -34.8337 | -23.676
14.3694 | -34.8342 | -23.6843
14.3755 | -34.8347 | -23.6926
14.3816 | -34.8352 | -23.7013
14.3877 | -34.8357 | -23.7095
14.3938 | -34.8362 | -23.7181
14.4 | -34.8367 | -23.7265
14.4061 | -34.8371 | -23.7348
14.4122 | -34.8376 | -23.7432
14.4183 | -34.8381 | -23.7516
14.4244 | -34.8386 | -23.7599
14.4305 | -34.8391 | -23.7682
14.4366 | -34.8396 | -23.7765
14.4427 | -34.8401 | -23.785
14.4488 | -34.8406 | -23.7935
14.4549 | -34.8411 | -23.8017
14.461 | -34.8416 | -23.81
14.4671 | -34.842 | -23.8184
14.4732 | -34.8425 | -23.8268
14.4793 | -34.843 | -23.8352
14.4854 | -34.8435 | -23.8435
14.4915 | -34.844 | -23.8516
14.4976 | -34.8445 | -23.8598
14.5037 | -34.845 | -23.8682
14.5098 | -34.8455 | -23.8766
14.5159 | -34.846 | -23.885
14.522 | -34.8465 | -23.8933
14.5281 | -34.847 | -23.9014
14.5342 | -34.8475 | -23.9099
14.5403 | -34.848 | -23.9181
14.5465 | -34.8484 | -23.9264
14.5526 | -34.8489 | -23.9348
14.5587 | -34.8494 | -23.9432
14.5648 | -34.8499 | -23.9513
14.5709 | -34.8504 | -23.9595
14.577 | -34.8509 | -23.9677
14.5831 | -34.8513 | -23.9761
14.5892 | -34.8518 | -23.9844
14.5953 | -34.8523 | -23.9925
14.6014 | -34.8528 | -24.001
14.6075 | -34.8533 | -24.0093
14.6136 | -34.8538 | -24.0177
14.6197 | -34.8542 | -24.0262
14.6258 | -34.8547 | -24.0347
14.6319 | -34.8552 | -24.0429
14.638 | -34.8557 | -24.0513
14.6441 | -34.8562 | -24.0597
14.6502 | -34.8567 | -24.0682
14.6563 | -34.8572 | -24.0764
14.6624 | -34.8577 | -24.0847
14.6685 | -34.8582 | -24.0928
14.6746 | -34.8586 | -24.1012
14.6807 | -34.8591 | -24.1093
14.6869 | -34.8596 | -24.1176
14.693 | -34.8602 | -24.1259
14.6991 | -34.8606 | -24.1342
14.7052 | -34.8612 | -24.1423
14.7113 | -34.8617 | -24.1506
14.7174 | -34.8622 | -24.1589
14.7235 | -34.8627 | -24.1669
14.7296 | -34.8631 | -24.1751
14.7357 | -34.8637 | -24.1833
14.7418 | -34.8642 | -24.1915
14.7479 | -34.8647 | -24.1996
14.754 | -34.8652 | -24.2081
14.7601 | -34.8657 | -24.2163
14.7662 | -34.8662 | -24.2245
14.7723 | -34.8667 | -24.2325
14.7784 | -34.8672 | -24.2409
14.7845 | -34.8677 | -24.2493
14.7906 | -34.8682 | -24.2577
14.7967 | -34.8687 | -24.2659
14.8028 | -34.8692 | -24.2742
14.8089 | -34.8697 | -24.2825
14.815 | -34.8702 | -24.2906
14.8211 | -34.8707 | -24.2988
14.8272 | -34.8712 | -24.307
14.8334 | -34.8717 | -24.3152
14.8395 | -34.8722 | -24.3235
14.8456 | -34.8727 | -24.3315
14.8517 | -34.8732 | -24.3399
14.8578 | -34.8737 | -24.3482
14.8639 | -34.8742 | -24.3566
14.87 | -34.8747 | -24.3649
14.8761 | -34.8752 | -24.3731
14.8822 | -34.8757 | -24.3814
14.8883 | -34.8762 | -24.3896
14.8944 | -34.8767 | -24.398
14.9005 | -34.8772 | -24.4063
14.9066 | -34.8777 | -24.4145
14.9127 | -34.8783 | -24.4226
14.9188 | -34.8787 | -24.4308
14.9249 | -34.8792 | -24.439
14.931 | -34.8797 | -24.4473
14.9371 | -34.8802 | -24.4554
14.9432 | -34.8808 | -24.4636
14.9493 | -34.8813 | -24.4717
14.9554 | -34.8818 | -24.4799
14.9615 | -34.8822 | -24.4882
14.9676 | -34.8827 | -24.4965
14.9738 | -34.8832 | -24.5047
14.9799 | -34.8837 | -24.513
14.986 | -34.8842 | -24.521
14.9921 | -34.8847 | -24.5293
14.9982 | -34.8852 | -24.5375
15.0043 | -34.8857 | -24.5456
15.0104 | -34.8862 | -24.5541
15.0165 | -34.8867 | -24.5625
15.0226 | -34.8872 | -24.5705
15.0287 | -34.8877 | -24.5787
15.0348 | -34.8882 | -24.5869
15.0409 | -34.8887 | -24.5951
15.047 | -34.8892 | -24.6032
15.0531 | -34.8898 | -24.6114
15.0592 | -34.8902 | -24.6197
15.0653 | -34.8908 | -24.628
15.0714 | -34.8913 | -24.636
15.0775 | -34.8918 | -24.6442
15.0836 | -34.8923 | -24.6523
15.0897 | -34.8928 | -24.6603
15.0958 | -34.8934 | -24.6685
15.1019 | -34.8939 | -24.6766
15.108 | -34.8944 | -24.6847
15.1141 | -34.8948 | -24.6929
15.1203 | -34.8953 | -24.7012
15.1264 | -34.8959 | -24.7092
15.1325 | -34.8964 | -24.7174
15.1386 | -34.8969 | -24.7255
15.1447 | -34.8974 | -24.7337
15.1508 | -34.898 | -24.7418
15.1569 | -34.8985 | -24.75
15.163 | -34.899 | -24.7579
15.1691 | -34.8995 | -24.766
15.1752 | -34.9001 | -24.7743
15.1813 | -34.9006 | -24.7825
15.1874 | -34.9011 | -24.7907
15.1935 | -34.9016 | -24.7987
15.1996 | -34.9021 | -24.8069
15.2057 | -34.9026 | -24.8151
15.2118 | -34.9031 | -24.8234
15.2179 | -34.9037 | -24.8316
15.224 | -34.9041 | -24.8398
15.2301 | -34.9047 | -24.8478
15.2362 | -34.9052 | -24.8559
15.2423 | -34.9057 | -24.8642
15.2484 | -34.9062 | -24.8724
15.2545 | -34.9067 | -24.8806
15.2607 | -34.9073 | -24.8888
15.2668 | -34.9078 | -24.8968
15.2729 | -34.9083 | -24.9049
15.279 | -34.9088 | -24.9134
15.2851 | -34.9094 | -24.9215
15.2912 | -34.9099 | -24.9296
15.2973 | -34.9104 | -24.9378
15.3034 | -34.9109 | -24.9457
15.3095 | -34.9115 | -24.9539
15.3156 | -34.912 | -24.962
15.3217 | -34.9125 | -24.9702
15.3278 | -34.913 | -24.9784
15.3339 | -34.9136 | -24.9864
15.34 | -34.9141 | -24.9944
15.3461 | -34.9146 | -25.0024
15.3522 | -34.9151 | -25.0107
15.3583 | -34.9156 | -25.0191
15.3644 | -34.9161 | -25.0271
15.3705 | -34.9167 | -25.0349
15.3766 | -34.9172 | -25.0431
15.3827 | -34.9177 | -25.0511
15.3888 | -34.9183 | -25.0591
15.3949 | -34.9188 | -25.0672
15.401 | -34.9193 | -25.0756
15.4072 | -34.9198 | -25.0834
15.4133 | -34.9203 | -25.0915
15.4194 | -34.9208 | -25.0996
15.4255 | -34.9213 | -25.1076
15.4316 | -34.9218 | -25.1158
15.4377 | -34.9224 | -25.1241
15.4438 | -34.9229 | -25.132
15.4499 | -34.9235 | -25.14
15.456 | -34.924 | -25.148
15.4621 | -34.9245 | -25.1562
15.4682 | -34.925 | -25.1643
15.4743 | -34.9256 | -25.1722
15.4804 | -34.9261 | -25.1805
15.4865 | -34.9265 | -25.1887
15.4926 | -34.9271 | -25.197
15.4987 | -34.9276 | -25.2052
15.5048 | -34.9281 | -25.2133
15.5109 | -34.9287 | -25.2212
15.517 | -34.9292 | -25.2294
15.5231 | -34.9297 | -25.2375
15.5292 | -34.9302 | -25.2456
15.5353 | -34.9308 | -25.2538
15.5414 | -34.9313 | -25.2619
15.5476 | -34.9318 | -25.2697
15.5537 | -34.9324 | -25.2778
15.5598 | -34.9329 | -25.2862
15.5659 | -34.9334 | -25.2942
15.572 | -34.9339 | -25.3023
15.5781 | -34.9344 | -25.3102
15.5842 | -34.935 | -25.3182
15.5903 | -34.9355 | -25.3261
15.5964 | -34.936 | -25.3342
15.6025 | -34.9366 | -25.3422
15.6086 | -34.9371 | -25.3503
15.6147 | -34.9376 | -25.3584
15.6208 | -34.9382 | -25.3662
15.6269 | -34.9387 | -25.3743
15.633 | -34.9392 | -25.3824
15.6391 | -34.9398 | -25.3905
15.6452 | -34.9403 | -25.3986
15.6513 | -34.9409 | -25.4064
15.6574 | -34.9414 | -25.4146
15.6635 | -34.9419 | -25.4226
15.6696 | -34.9425 | -25.4308
15.6757 | -34.943 | -25.4388
15.6818 | -34.9435 | -25.4468
15.688 | -34.944 | -25.4548
15.6941 | -34.9446 | -25.4626
15.7002 | -34.9451 | -25.4705
15.7063 | -34.9456 | -25.4784
15.7124 | -34.9461 | -25.4864
15.7185 | -34.9466 | -25.4945
15.7246 | -34.9472 | -25.5025
15.7307 | -34.9477 | -25.5104
15.7368 | -34.9482 | -25.5184
15.7429 | -34.9488 | -25.5263
15.749 | -34.9493 | -25.5346
15.7551 | -34.9498 | -25.5426
15.7612 | -34.9503 | -25.5504
15.7673 | -34.9509 | -25.5583
15.7734 | -34.9514 | -25.5662
15.7795 | -34.9519 | -25.5742
15.7856 | -34.9525 | -25.582
15.7917 | -34.953 | -25.5899
15.7978 | -34.9536 | -25.5981
15.8039 | -34.9541 | -25.6062
15.81 | -34.9546 | -25.6142
15.8161 | -34.9551 | -25.6222
15.8222 | -34.9557 | -25.6302
15.8283 | -34.9562 | -25.638
15.8345 | -34.9568 | -25.646
15.8406 | -34.9573 | -25.6538
15.8467 | -34.9578 | -25.662
15.8528 | -34.9584 | -25.67
15.8589 | -34.9589 | -25.6781
15.865 | -34.9595 | -25.6858
15.8711 | -34.96 | -25.6938
15.8772 | -34.9605 | -25.7016
15.8833 | -34.961 | -25.7096
15.8894 | -34.9616 | -25.7178
15.8955 | -34.9621 | -25.7257
15.9016 | -34.9626 | -25.7334
15.9077 | -34.9632 | -25.7414
15.9138 | -34.9637 | -25.7494
15.9199 | -34.9642 | -25.7572
15.926 | -34.9647 | -25.765
15.9321 | -34.9653 | -25.7728
15.9382 | -34.9658 | -25.7806
15.9443 | -34.9663 | -25.7886
15.9504 | -34.9669 | -25.7966
15.9565 | -34.9674 | -25.8044
15.9626 | -34.9679 | -25.8124
15.9687 | -34.9684 | -25.8201
15.9749 | -34.969 | -25.828
15.981 | -34.9695 | -25.836
15.9871 | -34.97 | -25.844
15.9932 | -34.9705 | -25.8519
15.9993 | -34.9711 | -25.8598
16.0054 | -34.9716 | -25.8675
16.0115 | -34.9722 | -25.8756
16.0176 | -34.9727 | -25.8836
16.0237 | -34.9732 | -25.8916
16.0298 | -34.9738 | -25.8995
16.0359 | -34.9743 | -25.9073
16.042 | -34.9748 | -25.9151
16.0481 | -34.9754 | -25.9231
16.0542 | -34.9759 | -25.931
16.0603 | -34.9764 | -25.9388
16.0664 | -34.977 | -25.9468
16.0725 | -34.9775 | -25.9545
16.0786 | -34.978 | -25.9625
16.0847 | -34.9786 | -25.9703
16.0908 | -34.9791 | -25.9783
16.0969 | -34.9796 | -25.9865
16.103 | -34.9801 | -25.9945
16.1091 | -34.9807 | -26.0023
16.1152 | -34.9813 | -26.0103
16.1214 | -34.9818 | -26.0182
16.1275 | -34.9823 | -26.0261
16.1336 | -34.9829 | -26.0339
16.1397 | -34.9834 | -26.042
16.1458 | -34.9839 | -26.0498
16.1519 | -34.9845 | -26.0578
16.158 | -34.985 | -26.0658
16.1641 | -34.9855 | -26.0737
16.1702 | -34.986 | -26.0816
16.1763 | -34.9866 | -26.0895
16.1824 | -34.9871 | -26.0972
16.1885 | -34.9876 | -26.1048
16.1946 | -34.9882 | -26.1125
16.2007 | -34.9887 | -26.1204
16.2068 | -34.9893 | -26.1283
16.2129 | -34.9898 | -26.1363
16.219 | -34.9903 | -26.144
16.2251 | -34.9909 | -26.1519
16.2312 | -34.9914 | -26.1597
16.2373 | -34.9919 | -26.1674
16.2434 | -34.9925 | -26.1753
16.2495 | -34.993 | -26.1829
16.2556 | -34.9936 | -26.191
16.2618 | -34.9941 | -26.1989
16.2679 | -34.9947 | -26.2068
16.274 | -34.9952 | -26.2146
16.2801 | -34.9957 | -26.2226
16.2862 | -34.9963 | -26.2303
16.2923 | -34.9969 | -26.2382
16.2984 | -34.9974 | -26.2461
16.3045 | -34.9979 | -26.2539
16.3106 | -34.9985 | -26.2618
16.3167 | -34.999 | -26.2696
16.3228 | -34.9995 | -26.2772
16.3289 | -35 | -26.2851
16.335 | -35.0006 | -26.2929
16.3411 | -35.0011 | -26.3006
16.3472 | -35.0017 | -26.3086
16.3533 | -35.0022 | -26.3165
16.3594 | -35.0028 | -26.324
16.3655 | -35.0033 | -26.3318
16.3716 | -35.0038 | -26.3396
16.3777 | -35.0043 | -26.3472
16.3838 | -35.0049 | -26.3551
16.3899 | -35.0054 | -26.363
16.396 | -35.006 | -26.3709
16.4021 | -35.0065 | -26.3787
16.4083 | -35.007 | -26.3863
16.4144 | -35.0075 | -26.3942
16.4205 | -35.0081 | -26.402
16.4266 | -35.0086 | -26.4096
16.4327 | -35.0091 | -26.4175
16.4388 | -35.0096 | -26.4254
16.4449 | -35.0102 | -26.4331
16.451 | -35.0107 | -26.441
16.4571 | -35.0112 | -26.4488
16.4632 | -35.0118 | -26.4562
16.4693 | -35.0123 | -26.4641
16.4754 | -35.0128 | -26.4718
16.4815 | -35.0134 | -26.4796
16.4876 | -35.0139 | -26.4872
16.4937 | -35.0144 | -26.495
16.4998 | -35.015 | -26.5025
16.5059 | -35.0155 | -26.5102
16.512 | -35.0161 | -26.5181
16.5181 | -35.0166 | -26.526
16.5242 | -35.0171 | -26.5339
16.5303 | -35.0177 | -26.5413
16.5364 | -35.0182 | -26.5491
16.5425 | -35.0187 | -26.5568
16.5487 | -35.0192 | -26.5645
16.5548 | -35.0198 | -26.5724
16.5609 | -35.0203 | -26.5802
16.567 | -35.0209 | -26.5878
16.5731 | -35.0214 | -26.5956
16.5792 | -35.0219 | -26.6034
16.5853 | -35.0225 | -26.6111
16.5914 | -35.023 | -26.619
16.5975 | -35.0235 | -26.6268
16.6036 | -35.0241 | -26.6345
16.6097 | -35.0246 | -26.6423
16.6158 | -35.0251 | -26.6501
16.6219 | -35.0257 | -26.6577
16.628 | -35.0262 | -26.6655
16.6341 | -35.0267 | -26.6732
16.6402 | -35.0272 | -26.6806
16.6463 | -35.0277 | -26.6883
16.6524 | -35.0283 | -26.6961
16.6585 | -35.0289 | -26.7038
16.6646 | -35.0294 | -26.7115
16.6707 | -35.0299 | -26.719
16.6768 | -35.0305 | -26.7268
16.6829 | -35.031 | -26.7345
16.689 | -35.0315 | -26.7423
16.6952 | -35.032 | -26.7499
16.7013 | -35.0326 | -26.7576
16.7074 | -35.0331 | -26.7651
16.7135 | -35.0337 | -26.7729
16.7196 | -35.0342 | -26.7806
16.7257 | -35.0348 | -26.7882
16.7318 | -35.0353 | -26.7958
16.7379 | -35.0358 | -26.8034
16.744 | -35.0363 | -26.811
16.7501 | -35.0369 | -26.8188
16.7562 | -35.0374 | -26.8264
16.7623 | -35.0379 | -26.834
16.7684 | -35.0385 | -26.8418
16.7745 | -35.039 | -26.8497
16.7806 | -35.0396 | -26.8569
16.7867 | -35.0401 | -26.8646
16.7928 | -35.0406 | -26.8722
16.7989 | -35.0411 | -26.88
16.805 | -35.0417 | -26.8877
16.8111 | -35.0422 | -26.8952
16.8172 | -35.0428 | -26.9029
16.8233 | -35.0433 | -26.9105
16.8294 | -35.0438 | -26.9182
16.8356 | -35.0443 | -26.9259
16.8417 | -35.0449 | -26.9336
16.8478 | -35.0454 | -26.9411
16.8539 | -35.0459 | -26.9486
16.86 | -35.0465 | -26.9564
16.8661 | -35.047 | -26.964
16.8722 | -35.0476 | -26.9717
16.8783 | -35.0481 | -26.9795
16.8844 | -35.0486 | -26.987
16.8905 | -35.0492 | -26.9946
16.8966 | -35.0497 | -27.0021
16.9027 | -35.0502 | -27.0097
16.9088 | -35.0508 | -27.0176
16.9149 | -35.0513 | -27.0254
16.921 | -35.0518 | -27.0328
16.9271 | -35.0523 | -27.0406
16.9332 | -35.0529 | -27.0484
16.9393 | -35.0534 | -27.0562
16.9454 | -35.0539 | -27.0637
16.9515 | -35.0545 | -27.0714
16.9576 | -35.055 | -27.079
16.9637 | -35.0555 | -27.0866
16.9698 | -35.0561 | -27.0941
16.9759 | -35.0566 | -27.1016
16.9821 | -35.0571 | -27.1094
16.9882 | -35.0576 | -27.1167
16.9943 | -35.0581 | -27.1242
17.0004 | -35.0587 | -27.1318
17.0065 | -35.0592 | -27.1396
17.0126 | -35.0597 | -27.1471
17.0187 | -35.0602 | -27.1549
17.0248 | -35.0608 | -27.1622
17.0309 | -35.0613 | -27.1698
17.037 | -35.0618 | -27.1774
17.0431 | -35.0623 | -27.185
17.0492 | -35.0628 | -27.1925
17.0553 | -35.0633 | -27.1999
17.0614 | -35.0638 | -27.2074
17.0675 | -35.0644 | -27.215
17.0736 | -35.0649 | -27.2227
17.0797 | -35.0654 | -27.2303
17.0858 | -35.0659 | -27.2379
17.0919 | -35.0665 | -27.2456
17.098 | -35.067 | -27.2531
17.1041 | -35.0675 | -27.2608
17.1102 | -35.0681 | -27.2683
17.1163 | -35.0687 | -27.2758
17.1225 | -35.0692 | -27.2832
17.1286 | -35.0697 | -27.2904
17.1347 | -35.0702 | -27.298
17.1408 | -35.0708 | -27.3057
17.1469 | -35.0713 | -27.3134
17.153 | -35.0718 | -27.3209
17.1591 | -35.0723 | -27.3285
17.1652 | -35.0729 | -27.3361
17.1713 | -35.0734 | -27.3435
17.1774 | -35.0739 | -27.3512
17.1835 | -35.0745 | -27.3587
17.1896 | -35.075 | -27.3662
17.1957 | -35.0755 | -27.3737
17.2018 | -35.076 | -27.381
17.2079 | -35.0766 | -27.3885
17.214 | -35.0771 | -27.3963
17.2201 | -35.0776 | -27.4039
17.2262 | -35.0781 | -27.4118
17.2323 | -35.0787 | -27.4193
17.2384 | -35.0792 | -27.4267
17.2445 | -35.0797 | -27.4342
17.2506 | -35.0803 | -27.4418
17.2567 | -35.0808 | -27.4495
17.2628 | -35.0813 | -27.4569
17.269 | -35.0818 | -27.4641
17.2751 | -35.0824 | -27.4718
17.2812 | -35.0829 | -27.4795
17.2873 | -35.0835 | -27.487
17.2934 | -35.084 | -27.4946
17.2995 | -35.0845 | -27.5022
17.3056 | -35.085 | -27.5097
17.3117 | -35.0856 | -27.5174
17.3178 | -35.0861 | -27.5249
17.3239 | -35.0866 | -27.5324
17.33 | -35.0871 | -27.54
17.3361 | -35.0876 | -27.5474
17.3422 | -35.0881 | -27.5546
17.3483 | -35.0887 | -27.5622
17.3544 | -35.0892 | -27.5697
17.3605 | -35.0897 | -27.5773
17.3666 | -35.0902 | -27.5848
17.3727 | -35.0908 | -27.5924
17.3788 | -35.0913 | -27.5996
17.3849 | -35.0918 | -27.6071
17.391 | -35.0923 | -27.6146
17.3971 | -35.0928 | -27.6221
17.4032 | -35.0934 | -27.6296
17.4094 | -35.0939 | -27.6368
17.4155 | -35.0944 | -27.6443
17.4216 | -35.0949 | -27.6519
17.4277 | -35.0954 | -27.6595
17.4338 | -35.096 | -27.6671
17.4399 | -35.0965 | -27.6748
17.446 | -35.097 | -27.682
17.4521 | -35.0975 | -27.6895
17.4582 | -35.0981 | -27.6969
17.4643 | -35.0986 | -27.7043
17.4704 | -35.0991 | -27.7117
17.4765 | -35.0996 | -27.7192
17.4826 | -35.1001 | -27.7265
17.4887 | -35.1006 | -27.7341
17.4948 | -35.1012 | -27.7417
17.5009 | -35.1017 | -27.7492
17.507 | -35.1022 | -27.7568
17.5131 | -35.1027 | -27.7643
17.5192 | -35.1032 | -27.7715
17.5253 | -35.1038 | -27.7788
17.5314 | -35.1043 | -27.7864
17.5375 | -35.1048 | -27.7938
17.5436 | -35.1053 | -27.8014
17.5497 | -35.1058 | -27.8089
17.5559 | -35.1063 | -27.816
17.562 | -35.1068 | -27.8233
17.5681 | -35.1073 | -27.8308
17.5742 | -35.1079 | -27.8383
17.5803 | -35.1084 | -27.8459
17.5864 | -35.1089 | -27.8532
17.5925 | -35.1094 | -27.8607
17.5986 | -35.11 | -27.8681
17.6047 | -35.1105 | -27.8758
17.6108 | -35.111 | -27.8832
17.6169 | -35.1115 | -27.8906
17.623 | -35.112 | -27.8978
17.6291 | -35.1125 | -27.905
17.6352 | -35.113 | -27.9125
17.6413 | -35.1136 | -27.92
17.6474 | -35.1141 | -27.9274
17.6535 | -35.1146 | -27.9349
17.6596 | -35.1151 | -27.9422
17.6657 | -35.1156 | -27.9498
17.6718 | -35.1162 | -27.9572
17.6779 | -35.1167 | -27.9647
17.684 | -35.1172 | -27.9722
17.6901 | -35.1177 | -27.9797
17.6963 | -35.1182 | -27.9869
17.7024 | -35.1187 | -27.9943
17.7085 | -35.1193 | -28.0016
17.7146 | -35.1198 | -28.0089
17.7207 | -35.1203 | -28.0162
17.7268 | -35.1208 | -28.0234
17.7329 | -35.1213 | -28.0309
17.739 | -35.1218 | -28.0384
17.7451 | -35.1223 | -28.0457
17.7512 | -35.1228 | -28.053
17.7573 | -35.1233 | -28.0606
17.7634 | -35.1238 | -28.0677
17.7695 | -35.1243 | -28.0751
17.7756 | -35.1249 | -28.0825
17.7817 | -35.1253 | -28.0899
17.7878 | -35.1259 | -28.0973
17.7939 | -35.1264 | -28.1047
17.8 | -35.1269 | -28.1119
17.8061 | -35.1274 | -28.1192
17.8122 | -35.128 | -28.1265
17.8183 | -35.1285 | -28.134
17.8244 | -35.129 | -28.1413
17.8305 | -35.1295 | -28.1487
17.8366 | -35.13 | -28.1557
17.8428 | -35.1305 | -28.1631
17.8489 | -35.131 | -28.1704
17.855 | -35.1315 | -28.1777
17.8611 | -35.132 | -28.1849
17.8672 | -35.1325 | -28.1922
17.8733 | -35.133 | -28.1996
17.8794 | -35.1335 | -28.207
17.8855 | -35.134 | -28.2144
17.8916 | -35.1345 | -28.2219
17.8977 | -35.135 | -28.2294
17.9038 | -35.1355 | -28.2365
17.9099 | -35.1361 | -28.2439
17.916 | -35.1366 | -28.2513
17.9221 | -35.1371 | -28.2587
17.9282 | -35.1376 | -28.2659
17.9343 | -35.1381 | -28.2733
17.9404 | -35.1386 | -28.2804
17.9465 | -35.1391 | -28.2877
17.9526 | -35.1396 | -28.2952
17.9587 | -35.1402 | -28.3026
17.9648 | -35.1407 | -28.3099
17.9709 | -35.1411 | -28.3172
17.977 | -35.1416 | -28.3244
17.9832 | -35.1422 | -28.3318
17.9893 | -35.1427 | -28.3391
17.9954 | -35.1432 | -28.3464
18.0015 | -35.1437 | -28.3539
18.0076 | -35.1442 | -28.361
18.0137 | -35.1447 | -28.3683
18.0198 | -35.1452 | -28.3757
18.0259 | -35.1457 | -28.383
18.032 | -35.1462 | -28.3904
18.0381 | -35.1467 | -28.3977
18.0442 | -35.1471 | -28.4048
18.0503 | -35.1477 | -28.4121
18.0564 | -35.1482 | -28.4194
18.0625 | -35.1487 | -28.4268
18.0686 | -35.1492 | -28.4341
18.0747 | -35.1496 | -28.4415
18.0808 | -35.1502 | -28.4486
18.0869 | -35.1507 | -28.4561
18.093 | -35.1512 | -28.4635
18.0991 | -35.1517 | -28.4709
18.1052 | -35.1522 | -28.4783
18.1113 | -35.1527 | -28.4854
18.1174 | -35.1532 | -28.4925
18.1236 | -35.1537 | -28.4997
18.1297 | -35.1542 | -28.507
18.1358 | -35.1547 | -28.5143
18.1419 | -35.1553 | -28.5216
18.148 | -35.1558 | -28.5288
18.1541 | -35.1563 | -28.5359
18.1602 | -35.1568 | -28.5433
18.1663 | -35.1573 | -28.5504
18.1724 | -35.1578 | -28.5578
18.1785 | -35.1583 | -28.565
18.1846 | -35.1588 | -28.5722
18.1907 | -35.1593 | -28.5794
18.1968 | -35.1598 | -28.5868
18.2029 | -35.1603 | -28.5943
18.209 | -35.1608 | -28.6016
18.2151 | -35.1613 | -28.609
18.2212 | -35.1618 | -28.6161
18.2273 | -35.1623 | -28.6235
18.2334 | -35.1628 | -28.6307
18.2395 | -35.1634 | -28.638
18.2456 | -35.1639 | -28.6453
18.2517 | -35.1644 | -28.6525
18.2578 | -35.1648 | -28.6597
18.2639 | -35.1654 | -28.667
18.2701 | -35.1659 | -28.6743
18.2762 | -35.1664 | -28.6814
18.2823 | -35.1669 | -28.6884
18.2884 | -35.1674 | -28.6956
18.2945 | -35.1679 | -28.7027
18.3006 | -35.1684 | -28.71
18.3067 | -35.1689 | -28.7171
18.3128 | -35.1694 | -28.7245
18.3189 | -35.1699 | -28.7317
18.325 | -35.1704 | -28.7388
18.3311 | -35.1709 | -28.7459
18.3372 | -35.1714 | -28.7531
18.3433 | -35.1719 | -28.7604
18.3494 | -35.1724 | -28.7676
18.3555 | -35.1729 | -28.7747
18.3616 | -35.1735 | -28.7817
18.3677 | -35.174 | -28.7892
18.3738 | -35.1745 | -28.7964
18.3799 | -35.175 | -28.8037
18.386 | -35.1755 | -28.811
18.3921 | -35.176 | -28.8185
18.3982 | -35.1765 | -28.8258
18.4043 | -35.177 | -28.833
18.4105 | -35.1775 | -28.8404
18.4166 | -35.178 | -28.8476
18.4227 | -35.1785 | -28.855
18.4288 | -35.179 | -28.8621
18.4349 | -35.1795 | -28.8689
18.441 | -35.18 | -28.876
18.4471 | -35.1805 | -28.8832
18.4532 | -35.181 | -28.8903
18.4593 | -35.1815 | -28.8975
18.4654 | -35.182 | -28.9043
18.4715 | -35.1825 | -28.9116
18.4776 | -35.1829 | -28.9191
18.4837 | -35.1835 | -28.9263
18.4898 | -35.1839 | -28.9337
18.4959 | -35.1844 | -28.941
18.502 | -35.185 | -28.9483
18.5081 | -35.1855 | -28.9555
18.5142 | -35.1859 | -28.9629
18.5203 | -35.1865 | -28.9701
18.5264 | -35.1869 | -28.9774
18.5325 | -35.1874 | -28.9846
18.5386 | -35.1879 | -28.9917
18.5447 | -35.1884 | -28.999
18.5508 | -35.1889 | -29.0061
18.557 | -35.1894 | -29.0134
18.5631 | -35.1899 | -29.0208
18.5692 | -35.1904 | -29.0281
18.5753 | -35.1909 | -29.035
18.5814 | -35.1914 | -29.0423
18.5875 | -35.1919 | -29.0493
18.5936 | -35.1924 | -29.0565
18.5997 | -35.1929 | -29.0637
18.6058 | -35.1934 | -29.0706
18.6119 | -35.1939 | -29.0778
18.618 | -35.1944 | -29.0849
18.6241 | -35.1949 | -29.0923
18.6302 | -35.1954 | -29.0996
18.6363 | -35.1959 | -29.1068
18.6424 | -35.1964 | -29.1139
18.6485 | -35.1969 | -29.1211
18.6546 | -35.1974 | -29.1284
18.6607 | -35.1979 | -29.1357
18.6668 | -35.1984 | -29.143
18.6729 | -35.1989 | -29.1504
18.679 | -35.1994 | -29.1575
18.6851 | -35.1999 | -29.1646
18.6912 | -35.2004 | -29.1717
18.6974 | -35.2009 | -29.1789
18.7035 | -35.2014 | -29.1861
18.7096 | -35.2019 | -29.1934
18.7157 | -35.2024 | -29.2006
18.7218 | -35.2029 | -29.2077
18.7279 | -35.2034 | -29.215
18.734 | -35.2039 | -29.2221
18.7401 | -35.2044 | -29.2293
18.7462 | -35.2049 | -29.2363
18.7523 | -35.2054 | -29.2434
18.7584 | -35.2059 | -29.2506
18.7645 | -35.2064 | -29.2578
18.7706 | -35.2069 | -29.265
18.7767 | -35.2074 | -29.2723
18.7828 | -35.2079 | -29.2794
18.7889 | -35.2085 | -29.2868
18.795 | -35.209 | -29.2941
18.8011 | -35.2095 | -29.3013
18.8072 | -35.21 | -29.3086
18.8133 | -35.2105 | -29.3158
18.8194 | -35.211 | -29.3228
18.8255 | -35.2115 | -29.3299
18.8316 | -35.212 | -29.3373
18.8377 | -35.2125 | -29.3444
18.8439 | -35.213 | -29.3515
18.85 | -35.2135 | -29.3588
18.8561 | -35.214 | -29.3659
18.8622 | -35.2145 | -29.373
18.8683 | -35.215 | -29.38
18.8744 | -35.2155 | -29.3872
18.8805 | -35.216 | -29.3945
18.8866 | -35.2165 | -29.4018
18.8927 | -35.217 | -29.4086
18.8988 | -35.2175 | -29.4159
18.9049 | -35.218 | -29.423
18.911 | -35.2185 | -29.4302
18.9171 | -35.219 | -29.4373
18.9232 | -35.2195 | -29.4444
18.9293 | -35.22 | -29.4515
18.9354 | -35.2206 | -29.4587
18.9415 | -35.2211 | -29.466
18.9476 | -35.2216 | -29.473
18.9537 | -35.2221 | -29.4803
18.9598 | -35.2226 | -29.4871
18.9659 | -35.223 | -29.4943
18.972 | -35.2236 | -29.5012
18.9781 | -35.224 | -29.5085
18.9843 | -35.2245 | -29.5157
18.9904 | -35.225 | -29.5228
18.9965 | -35.2256 | -29.5297
19.0026 | -35.2261 | -29.5369
19.0087 | -35.2266 | -29.5441
19.0148 | -35.2271 | -29.551
19.0209 | -35.2276 | -29.5581
19.027 | -35.2281 | -29.5651
19.0331 | -35.2286 | -29.5721
19.0392 | -35.229 | -29.5793
19.0453 | -35.2295 | -29.5865
19.0514 | -35.23 | -29.5935
19.0575 | -35.2305 | -29.6007
19.0636 | -35.231 | -29.6075
19.0697 | -35.2316 | -29.6147
19.0758 | -35.232 | -29.6217
19.0819 | -35.2325 | -29.6289
19.088 | -35.2331 | -29.6359
19.0941 | -35.2335 | -29.6432
19.1002 | -35.234 | -29.6504
19.1063 | -35.2344 | -29.6575
19.1124 | -35.2349 | -29.6644
19.1185 | -35.2354 | -29.6717
19.1246 | -35.2359 | -29.6787
19.1308 | -35.2364 | -29.6859
19.1369 | -35.2369 | -29.6928
19.143 | -35.2374 | -29.7
19.1491 | -35.2379 | -29.7072
19.1552 | -35.2384 | -29.7145
19.1613 | -35.2389 | -29.7215
19.1674 | -35.2394 | -29.7286
19.1735 | -35.2399 | -29.7358
19.1796 | -35.2404 | -29.7428
19.1857 | -35.2409 | -29.7499
19.1918 | -35.2414 | -29.757
19.1979 | -35.2419 | -29.7642
19.204 | -35.2424 | -29.7712
19.2101 | -35.2429 | -29.7782
19.2162 | -35.2434 | -29.7852
19.2223 | -35.2439 | -29.7922
19.2284 | -35.2444 | -29.7993
19.2345 | -35.2449 | -29.8064
19.2406 | -35.2454 | -29.8134
19.2467 | -35.2459 | -29.8208
19.2528 | -35.2464 | -29.8282
19.2589 | -35.2469 | -29.8352
19.265 | -35.2473 | -29.8422
19.2712 | -35.2478 | -29.8493
19.2773 | -35.2484 | -29.8562
19.2834 | -35.2489 | -29.8633
19.2895 | -35.2494 | -29.8706
19.2956 | -35.2499 | -29.8778
19.3017 | -35.2504 | -29.8849
19.3078 | -35.2509 | -29.8919
19.3139 | -35.2514 | -29.8989
19.32 | -35.2519 | -29.9059
19.3261 | -35.2523 | -29.9131
19.3322 | -35.2528 | -29.92
19.3383 | -35.2533 | -29.9269
19.3444 | -35.2538 | -29.9338
19.3505 | -35.2543 | -29.9408
19.3566 | -35.2548 | -29.9478
19.3627 | -35.2553 | -29.9549
19.3688 | -35.2558 | -29.9619
19.3749 | -35.2563 | -29.9689
19.381 | -35.2568 | -29.9758
19.3871 | -35.2573 | -29.9826
19.3932 | -35.2577 | -29.9898
19.3993 | -35.2582 | -29.9968
19.4054 | -35.2587 | -30.0039
19.4115 | -35.2592 | -30.0108
19.4177 | -35.2598 | -30.0177
19.4238 | -35.2603 | -30.0247
19.4299 | -35.2608 | -30.0317
19.436 | -35.2613 | -30.039
19.4421 | -35.2618 | -30.046
19.4482 | -35.2623 | -30.0533
19.4543 | -35.2628 | -30.0603
19.4604 | -35.2632 | -30.0672
19.4665 | -35.2637 | -30.074
19.4726 | -35.2642 | -30.081
19.4787 | -35.2647 | -30.088
19.4848 | -35.2652 | -30.0952
19.4909 | -35.2657 | -30.1019
19.497 | -35.2662 | -30.1088
19.5031 | -35.2667 | -30.1159
19.5092 | -35.2672 | -30.123
19.5153 | -35.2677 | -30.13
19.5214 | -35.2682 | -30.1371
19.5275 | -35.2687 | -30.1441
19.5336 | -35.2691 | -30.1511
19.5397 | -35.2696 | -30.1579
19.5458 | -35.2701 | -30.1649
19.5519 | -35.2706 | -30.1721
19.5581 | -35.2711 | -30.1792
19.5642 | -35.2716 | -30.1863
19.5703 | -35.2721 | -30.1936
19.5764 | -35.2726 | -30.2005
19.5825 | -35.2731 | -30.2075
19.5886 | -35.2736 | -30.2148
19.5947 | -35.2742 | -30.2217
19.6008 | -35.2747 | -30.2286
19.6069 | -35.2752 | -30.2358
19.613 | -35.2757 | -30.2428
19.6191 | -35.2762 | -30.25
19.6252 | -35.2767 | -30.257
19.6313 | -35.2772 | -30.2639
19.6374 | -35.2777 | -30.2712
19.6435 | -35.2782 | -30.2782
19.6496 | -35.2787 | -30.2854
19.6557 | -35.2792 | -30.2924
19.6618 | -35.2797 | -30.2991
19.6679 | -35.2802 | -30.3061
19.674 | -35.2807 | -30.3131
19.6801 | -35.2812 | -30.3201
19.6862 | -35.2817 | -30.3271
19.6923 | -35.2822 | -30.3339
19.6984 | -35.2827 | -30.3409
19.7046 | -35.2832 | -30.348
19.7107 | -35.2837 | -30.3548
19.7168 | -35.2843 | -30.3618
19.7229 | -35.2848 | -30.3687
19.729 | -35.2853 | -30.3757
19.7351 | -35.2858 | -30.3828
19.7412 | -35.2863 | -30.3899
19.7473 | -35.2868 | -30.397
19.7534 | -35.2873 | -30.4041
19.7595 | -35.2878 | -30.4112
19.7656 | -35.2883 | -30.4184
19.7717 | -35.2888 | -30.4248
19.7778 | -35.2893 | -30.4317
19.7839 | -35.2898 | -30.4386
19.79 | -35.2903 | -30.4454
19.7961 | -35.2908 | -30.4523
19.8022 | -35.2913 | -30.4589
19.8083 | -35.2918 | -30.4657
19.8144 | -35.2923 | -30.4728
19.8205 | -35.2928 | -30.4795
19.8266 | -35.2934 | -30.4864
19.8327 | -35.2939 | -30.4932
19.8388 | -35.2944 | -30.5
19.845 | -35.2949 | -30.5067
19.8511 | -35.2954 | -30.5141
19.8572 | -35.2959 | -30.521
19.8633 | -35.2964 | -30.528
19.8694 | -35.2969 | -30.535
19.8755 | -35.2974 | -30.5418
19.8816 | -35.2979 | -30.5485
19.8877 | -35.2984 | -30.5556
19.8938 | -35.2989 | -30.5625
19.8999 | -35.2994 | -30.5695
19.906 | -35.2999 | -30.5764
19.9121 | -35.3004 | -30.5833
19.9182 | -35.3009 | -30.5903
19.9243 | -35.3014 | -30.5971
19.9304 | -35.3019 | -30.604
19.9365 | -35.3025 | -30.611
19.9426 | -35.303 | -30.6179
19.9487 | -35.3035 | -30.6247
19.9548 | -35.304 | -30.6315
19.9609 | -35.3045 | -30.6384
19.967 | -35.305 | -30.6454
19.9731 | -35.3055 | -30.6521
19.9792 | -35.306 | -30.6591
19.9853 | -35.3065 | -30.6658
19.9915 | -35.307 | -30.6725
19.9976 | -35.3075 | -30.6796
20.0037 | -35.308 | -30.6864
20.0098 | -35.3085 | -30.6933
20.0159 | -35.309 | -30.7
20.022 | -35.3095 | -30.7071
20.0281 | -35.31 | -30.714
20.0342 | -35.3105 | -30.721
20.0403 | -35.3111 | -30.7281
20.0464 | -35.3116 | -30.7348
20.0525 | -35.3121 | -30.7412
20.0586 | -35.3125 | -30.7482
20.0647 | -35.313 | -30.7553
20.0708 | -35.3135 | -30.7626
20.0769 | -35.314 | -30.7693
20.083 | -35.3146 | -30.7762
20.0891 | -35.3151 | -30.7831
20.0952 | -35.3155 | -30.7901
20.1013 | -35.316 | -30.7969
20.1074 | -35.3166 | -30.8039
20.1135 | -35.3171 | -30.8111
20.1196 | -35.3176 | -30.8177
20.1257 | -35.318 | -30.8249
20.1319 | -35.3186 | -30.8317
20.138 | -35.3191 | -30.8381
20.1441 | -35.3196 | -30.845
20.1502 | -35.3201 | -30.8516
20.1563 | -35.3206 | -30.8582
20.1624 | -35.3211 | -30.865
20.1685 | -35.3216 | -30.8719
20.1746 | -35.322 | -30.8793
20.1807 | -35.3225 | -30.886
20.1868 | -35.323 | -30.8928
20.1929 | -35.3235 | -30.8997
20.199 | -35.3241 | -30.9065
20.2051 | -35.3245 | -30.9137
20.2112 | -35.325 | -30.9207
20.2173 | -35.3255 | -30.9275
20.2234 | -35.326 | -30.9342
20.2295 | -35.3265 | -30.9409
20.2356 | -35.3271 | -30.9478
20.2417 | -35.3276 | -30.9544
20.2478 | -35.3281 | -30.9612
20.2539 | -35.3286 | -30.9679
20.26 | -35.3291 | -30.9746
20.2661 | -35.3296 | -30.9818
20.2723 | -35.3301 | -30.9886
20.2784 | -35.3306 | -30.9958
20.2845 | -35.3312 | -31.0028
20.2906 | -35.3316 | -31.0097
20.2967 | -35.3321 | -31.0164
20.3028 | -35.3326 | -31.0232
20.3089 | -35.3331 | -31.0302
20.315 | -35.3336 | -31.0367
20.3211 | -35.3341 | -31.0436
20.3272 | -35.3346 | -31.0508
20.3333 | -35.3351 | -31.0577
20.3394 | -35.3357 | -31.0644
20.3455 | -35.3362 | -31.0712
20.3516 | -35.3367 | -31.0778
20.3577 | -35.3371 | -31.0847
20.3638 | -35.3377 | -31.0915
20.3699 | -35.3382 | -31.0985
20.376 | -35.3387 | -31.1052
20.3821 | -35.3392 | -31.1119
20.3882 | -35.3397 | -31.1185
20.3943 | -35.3402 | -31.1255
20.4004 | -35.3407 | -31.1324
20.4065 | -35.3413 | -31.1394
20.4126 | -35.3417 | -31.1461
20.4188 | -35.3422 | -31.1534
20.4249 | -35.3428 | -31.1603
20.431 | -35.3433 | -31.1669
20.4371 | -35.3439 | -31.1736
20.4432 | -35.3444 | -31.1807
20.4493 | -35.3448 | -31.1876
20.4554 | -35.3454 | -31.1945
20.4615 | -35.3459 | -31.2014
20.4676 | -35.3464 | -31.2085
20.4737 | -35.347 | -31.2155
20.4798 | -35.3475 | -31.2224
20.4859 | -35.348 | -31.2294
20.492 | -35.3486 | -31.236
20.4981 | -35.3491 | -31.2427
20.5042 | -35.3496 | -31.2494
20.5103 | -35.3501 | -31.2561
20.5164 | -35.3506 | -31.2633
20.5225 | -35.3511 | -31.2703
20.5286 | -35.3517 | -31.277
20.5347 | -35.3522 | -31.284
20.5408 | -35.3527 | -31.2909
20.5469 | -35.3532 | -31.2981
20.553 | -35.3537 | -31.305
20.5592 | -35.3543 | -31.3117
20.5653 | -35.3548 | -31.3183
20.5714 | -35.3553 | -31.3248
20.5775 | -35.3559 | -31.3317
20.5836 | -35.3563 | -31.3384
20.5897 | -35.3569 | -31.3454
20.5958 | -35.3574 | -31.3524
20.6019 | -35.3579 | -31.3591
20.608 | -35.3584 | -31.366
20.6141 | -35.3589 | -31.373
20.6202 | -35.3594 | -31.3795
20.6263 | -35.3599 | -31.3861
20.6324 | -35.3605 | -31.3926
20.6385 | -35.361 | -31.3995
20.6446 | -35.3615 | -31.4063
20.6507 | -35.362 | -31.4133
20.6568 | -35.3625 | -31.4202
20.6629 | -35.3631 | -31.4272
20.669 | -35.3636 | -31.4343
20.6751 | -35.3641 | -31.4413
20.6812 | -35.3647 | -31.4479
20.6873 | -35.3651 | -31.4544
20.6934 | -35.3656 | -31.4614
20.6995 | -35.3662 | -31.468
20.7057 | -35.3667 | -31.4748
20.7118 | -35.3673 | -31.4818
20.7179 | -35.3678 | -31.4884
20.724 | -35.3683 | -31.4953
20.7301 | -35.3688 | -31.5016
20.7362 | -35.3693 | -31.5082
20.7423 | -35.3698 | -31.515
20.7484 | -35.3703 | -31.5216
20.7545 | -35.3709 | -31.5281
20.7606 | -35.3715 | -31.5345
20.7667 | -35.372 | -31.5413
20.7728 | -35.3725 | -31.5479
20.7789 | -35.373 | -31.5551
20.785 | -35.3735 | -31.5618
20.7911 | -35.374 | -31.5683
20.7972 | -35.3745 | -31.575
20.8033 | -35.375 | -31.5817
20.8094 | -35.3755 | -31.5882
20.8155 | -35.376 | -31.5948
20.8216 | -35.3765 | -31.602
20.8277 | -35.377 | -31.6087
20.8338 | -35.3775 | -31.6156
20.8399 | -35.378 | -31.6223
20.8461 | -35.3786 | -31.6289
20.8522 | -35.3791 | -31.6355
20.8583 | -35.3797 | -31.6427
20.8644 | -35.3802 | -31.6495
20.8705 | -35.3807 | -31.6562
20.8766 | -35.3811 | -31.6631
20.8827 | -35.3816 | -31.6699
20.8888 | -35.3821 | -31.6768
20.8949 | -35.3826 | -31.6829
20.901 | -35.3831 | -31.6897
20.9071 | -35.3836 | -31.6965
20.9132 | -35.3842 | -31.7037
20.9193 | -35.3847 | -31.7104
20.9254 | -35.3852 | -31.7171
20.9315 | -35.3857 | -31.7237
20.9376 | -35.3862 | -31.7303
20.9437 | -35.3866 | -31.737
20.9498 | -35.3871 | -31.7435
20.9559 | -35.3876 | -31.75
20.962 | -35.3882 | -31.7568
20.9681 | -35.3887 | -31.7636
20.9742 | -35.3893 | -31.7702
20.9803 | -35.3898 | -31.7773
20.9864 | -35.3903 | -31.7839
20.9926 | -35.3908 | -31.791
20.9987 | -35.3913 | -31.7974
21.0048 | -35.3918 | -31.8043
21.0109 | -35.3923 | -31.811
21.017 | -35.3928 | -31.8178
21.0231 | -35.3933 | -31.8247
21.0292 | -35.3939 | -31.831
21.0353 | -35.3943 | -31.8371
21.0414 | -35.3949 | -31.844
21.0475 | -35.3953 | -31.8509
21.0536 | -35.3958 | -31.8578
21.0597 | -35.3963 | -31.8646
21.0658 | -35.3968 | -31.8714
21.0719 | -35.3974 | -31.8784
21.078 | -35.3979 | -31.885
21.0841 | -35.3984 | -31.8916
21.0902 | -35.3989 | -31.8985
21.0963 | -35.3994 | -31.9054
21.1024 | -35.4001 | -31.9116
21.1085 | -35.4007 | -31.9176
21.1146 | -35.4012 | -31.9242
21.1207 | -35.4017 | -31.931
21.1268 | -35.4022 | -31.9375
21.133 | -35.4027 | -31.9441
21.1391 | -35.4032 | -31.9507
21.1452 | -35.4036 | -31.9574
21.1513 | -35.4042 | -31.9639
21.1574 | -35.4047 | -31.9709
21.1635 | -35.4052 | -31.9773
21.1696 | -35.4058 | -31.9841
21.1757 | -35.4064 | -31.9903
21.1818 | -35.4069 | -31.9969
21.1879 | -35.4074 | -32.0033
21.194 | -35.4079 | -32.0104
21.2001 | -35.4085 | -32.0173
21.2062 | -35.409 | -32.024
21.2123 | -35.4095 | -32.0303
21.2184 | -35.41 | -32.037
21.2245 | -35.4105 | -32.0434
21.2306 | -35.4111 | -32.0502
21.2367 | -35.4116 | -32.0569
21.2428 | -35.4121 | -32.063
21.2489 | -35.4126 | -32.0696
21.255 | -35.4132 | -32.0765
21.2611 | -35.4137 | -32.083
21.2672 | -35.4143 | -32.0899
21.2733 | -35.4148 | -32.0967
21.2795 | -35.4153 | -32.1031
21.2856 | -35.4158 | -32.1102
21.2917 | -35.4164 | -32.1169
21.2978 | -35.4169 | -32.1239
21.3039 | -35.4174 | -32.1308
21.31 | -35.4179 | -32.1369
21.3161 | -35.4184 | -32.1434
21.3222 | -35.4189 | -32.1498
21.3283 | -35.4194 | -32.1564
21.3344 | -35.4199 | -32.1628
21.3405 | -35.4205 | -32.1698
21.3466 | -35.421 | -32.1768
21.3527 | -35.4216 | -32.1831
21.3588 | -35.4221 | -32.1897
21.3649 | -35.4227 | -32.1966
21.371 | -35.4232 | -32.2033
21.3771 | -35.4237 | -32.2101
21.3832 | -35.4242 | -32.2164
21.3893 | -35.4247 | -32.223
21.3954 | -35.4252 | -32.2297
21.4015 | -35.4257 | -32.2358
21.4076 | -35.4263 | -32.2424
21.4137 | -35.4269 | -32.2493
21.4199 | -35.4274 | -32.2562
21.426 | -35.4279 | -32.2627
21.4321 | -35.4284 | -32.2696
21.4382 | -35.429 | -32.2764
21.4443 | -35.4295 | -32.2829
21.4504 | -35.43 | -32.2896
21.4565 | -35.4306 | -32.2956
21.4626 | -35.431 | -32.3025
21.4687 | -35.4316 | -32.3093
21.4748 | -35.4321 | -32.3158
21.4809 | -35.4326 | -32.3221
21.487 | -35.4331 | -32.3287
21.4931 | -35.4337 | -32.3351
21.4992 | -35.4342 | -32.3418
21.5053 | -35.4347 | -32.3486
21.5114 | -35.4352 | -32.3554
21.5175 | -35.4357 | -32.3617
21.5236 | -35.4361 | -32.3686
21.5297 | -35.4367 | -32.3744
21.5358 | -35.4372 | -32.3812
21.5419 | -35.4377 | -32.388
21.548 | -35.4382 | -32.3945
21.5541 | -35.4387 | -32.4013
21.5602 | -35.4392 | -32.4077
21.5664 | -35.4397 | -32.4142
21.5725 | -35.4402 | -32.421
21.5786 | -35.4407 | -32.4276
21.5847 | -35.4412 | -32.4341
21.5908 | -35.4417 | -32.4403
21.5969 | -35.4423 | -32.4463
21.603 | -35.4428 | -32.4528
21.6091 | -35.4433 | -32.4595
21.6152 | -35.4438 | -32.4662
21.6213 | -35.4444 | -32.4726
21.6274 | -35.4449 | -32.4789
21.6335 | -35.4454 | -32.4857
21.6396 | -35.4459 | -32.4924
21.6457 | -35.4463 | -32.499
21.6518 | -35.4469 | -32.5054
21.6579 | -35.4474 | -32.512
21.664 | -35.448 | -32.5185
21.6701 | -35.4486 | -32.5254
21.6762 | -35.4492 | -32.5322
21.6823 | -35.4497 | -32.5392
21.6884 | -35.4502 | -32.5452
21.6945 | -35.4508 | -32.5516
21.7006 | -35.4513 | -32.5585
21.7068 | -35.4518 | -32.5647
21.7129 | -35.4523 | -32.5714
21.719 | -35.4528 | -32.5779
21.7251 | -35.4532 | -32.5846
21.7312 | -35.4538 | -32.5915
21.7373 | -35.4543 | -32.5981
21.7434 | -35.4548 | -32.6047
21.7495 | -35.4553 | -32.611
21.7556 | -35.4558 | -32.6176
21.7617 | -35.4563 | -32.6245
21.7678 | -35.4568 | -32.631
21.7739 | -35.4574 | -32.6374
21.78 | -35.4579 | -32.644
21.7861 | -35.4585 | -32.6508
21.7922 | -35.459 | -32.6575
21.7983 | -35.4595 | -32.6641
21.8044 | -35.46 | -32.6704
21.8105 | -35.4606 | -32.6771
21.8166 | -35.4611 | -32.6837
21.8227 | -35.4616 | -32.6901
21.8288 | -35.4622 | -32.6972
21.8349 | -35.4627 | -32.7039
21.841 | -35.4632 | -32.7107
21.8471 | -35.4637 | -32.7169
21.8533 | -35.4643 | -32.7235
21.8594 | -35.4648 | -32.73
21.8655 | -35.4653 | -32.7366
21.8716 | -35.4658 | -32.7434
21.8777 | -35.4663 | -32.7502
21.8838 | -35.4668 | -32.7567
21.8899 | -35.4673 | -32.7629
21.896 | -35.4679 | -32.7697
21.9021 | -35.4684 | -32.7764
21.9082 | -35.4689 | -32.783
21.9143 | -35.4695 | -32.7891
21.9204 | -35.47 | -32.7954
21.9265 | -35.4704 | -32.8018
21.9326 | -35.4709 | -32.8086
21.9387 | -35.4714 | -32.8154
21.9448 | -35.4719 | -32.8218
21.9509 | -35.4724 | -32.8282
21.957 | -35.4729 | -32.835
21.9631 | -35.4734 | -32.8416
21.9692 | -35.4739 | -32.8482
21.9753 | -35.4745 | -32.8544
21.9814 | -35.475 | -32.8606
21.9875 | -35.4755 | -32.8668
21.9937 | -35.476 | -32.8735
21.9998 | -35.4765 | -32.8803
22.0059 | -35.477 | -32.887
22.012 | -35.4775 | -32.893
22.0181 | -35.478 | -32.8991
22.0242 | -35.4784 | -32.9054
22.0303 | -35.4789 | -32.9119
22.0364 | -35.4794 | -32.9183
22.0425 | -35.4799 | -32.9248
22.0486 | -35.4805 | -32.9315
22.0547 | -35.481 | -32.9377
22.0608 | -35.4816 | -32.9441
22.0669 | -35.4821 | -32.9506
22.073 | -35.4826 | -32.9567
22.0791 | -35.4832 | -32.9633
22.0852 | -35.4837 | -32.9697
22.0913 | -35.4843 | -32.9756
22.0974 | -35.4848 | -32.982
22.1035 | -35.4853 | -32.9883
22.1096 | -35.4859 | -32.9944
22.1157 | -35.4864 | -33.0008
22.1218 | -35.487 | -33.0073
22.1279 | -35.4875 | -33.0138
22.134 | -35.488 | -33.0205
22.1402 | -35.4885 | -33.0265
22.1463 | -35.489 | -33.033
22.1524 | -35.4895 | -33.0394
22.1585 | -35.4901 | -33.0459
22.1646 | -35.4906 | -33.0519
22.1707 | -35.4912 | -33.0586
22.1768 | -35.4917 | -33.0654
22.1829 | -35.4922 | -33.072
22.189 | -35.4927 | -33.0786
22.1951 | -35.4933 | -33.0847
22.2012 | -35.4937 | -33.0911
22.2073 | -35.4943 | -33.0977
22.2134 | -35.4948 | -33.104
22.2195 | -35.4953 | -33.1105
22.2256 | -35.4958 | -33.1172
22.2317 | -35.4963 | -33.1237
22.2378 | -35.4968 | -33.1304
22.2439 | -35.4974 | -33.1369
22.25 | -35.4979 | -33.1432
22.2561 | -35.4984 | -33.1493
22.2622 | -35.4989 | -33.1555
22.2683 | -35.4994 | -33.1617
22.2744 | -35.4999 | -33.1683
22.2806 | -35.5005 | -33.175
22.2867 | -35.501 | -33.1811
22.2928 | -35.5015 | -33.1876
22.2989 | -35.502 | -33.1941
22.305 | -35.5025 | -33.2003
22.3111 | -35.503 | -33.2069
22.3172 | -35.5035 | -33.2132
22.3233 | -35.504 | -33.2195
22.3294 | -35.5046 | -33.2263
22.3355 | -35.5052 | -33.2324
22.3416 | -35.5057 | -33.2387
22.3477 | -35.5062 | -33.2454
22.3538 | -35.5067 | -33.2515
22.3599 | -35.5073 | -33.2581
22.366 | -35.5078 | -33.2646
22.3721 | -35.5083 | -33.2705
22.3782 | -35.5088 | -33.2766
22.3843 | -35.5094 | -33.2833
22.3904 | -35.5099 | -33.2892
22.3965 | -35.5103 | -33.2958
22.4026 | -35.5108 | -33.3024
22.4087 | -35.5114 | -33.3085
22.4148 | -35.5119 | -33.3149
22.4209 | -35.5124 | -33.3214
22.4271 | -35.5129 | -33.328
22.4332 | -35.5134 | -33.3342
22.4393 | -35.5139 | -33.3408
22.4454 | -35.5144 | -33.3468
22.4515 | -35.515 | -33.3532
22.4576 | -35.5155 | -33.3594
22.4637 | -35.5161 | -33.3661
22.4698 | -35.5166 | -33.3726
22.4759 | -35.5171 | -33.3788
22.482 | -35.5175 | -33.3854
22.4881 | -35.5181 | -33.3919
22.4942 | -35.5185 | -33.3982
22.5003 | -35.519 | -33.4048
22.5064 | -35.5196 | -33.4111
22.5125 | -35.52 | -33.4175
22.5186 | -35.5206 | -33.4238
22.5247 | -35.521 | -33.4304
22.5308 | -35.5214 | -33.4368
22.5369 | -35.5219 | -33.4433
22.543 | -35.5224 | -33.4499
22.5491 | -35.5229 | -33.4563
22.5552 | -35.5234 | -33.4623
22.5613 | -35.524 | -33.4687
22.5675 | -35.5245 | -33.4749
22.5736 | -35.525 | -33.4811
22.5797 | -35.5255 | -33.4876
22.5858 | -35.5261 | -33.4937
22.5919 | -35.5266 | -33.4999
22.598 | -35.5271 | -33.5065
22.6041 | -35.5276 | -33.5126
22.6102 | -35.5282 | -33.519
22.6163 | -35.5286 | -33.525
22.6224 | -35.5291 | -33.5315
22.6285 | -35.5296 | -33.538
22.6346 | -35.5301 | -33.5445
22.6407 | -35.5306 | -33.5509
22.6468 | -35.5312 | -33.5569
22.6529 | -35.5317 | -33.563
22.659 | -35.5322 | -33.5692
22.6651 | -35.5326 | -33.5756
22.6712 | -35.5331 | -33.5821
22.6773 | -35.5336 | -33.5885
22.6834 | -35.5341 | -33.595
22.6895 | -35.5346 | -33.6012
22.6956 | -35.5351 | -33.6075
22.7017 | -35.5356 | -33.6138
22.7079 | -35.5362 | -33.6202
22.714 | -35.5367 | -33.6264
22.7201 | -35.5372 | -33.633
22.7262 | -35.5378 | -33.6389
22.7323 | -35.5383 | -33.6449
22.7384 | -35.5388 | -33.6509
22.7445 | -35.5393 | -33.6574
22.7506 | -35.5398 | -33.6635
22.7567 | -35.5403 | -33.6699
22.7628 | -35.5409 | -33.676
22.7689 | -35.5414 | -33.6823
22.775 | -35.5419 | -33.6887
22.7811 | -35.5424 | -33.6951
22.7872 | -35.5429 | -33.7015
22.7933 | -35.5434 | -33.7078
22.7994 | -35.5439 | -33.7141
22.8055 | -35.5445 | -33.7207
22.8116 | -35.545 | -33.7269
22.8177 | -35.5456 | -33.7332
22.8238 | -35.546 | -33.7395
22.8299 | -35.5465 | -33.7456
22.836 | -35.547 | -33.7524
22.8421 | -35.5475 | -33.7588
22.8482 | -35.5481 | -33.7649
22.8544 | -35.5487 | -33.7709
22.8605 | -35.5492 | -33.7769
22.8666 | -35.5497 | -33.7833
22.8727 | -35.5502 | -33.7897
22.8788 | -35.5506 | -33.7959
22.8849 | -35.5512 | -33.8018
22.891 | -35.5517 | -33.8079
22.8971 | -35.5521 | -33.8145
22.9032 | -35.5526 | -33.8204
22.9093 | -35.5531 | -33.8264
22.9154 | -35.5536 | -33.8326
22.9215 | -35.5541 | -33.839
22.9276 | -35.5545 | -33.8453
22.9337 | -35.555 | -33.8513
22.9398 | -35.5555 | -33.8576
22.9459 | -35.5561 | -33.8638
22.952 | -35.5566 | -33.8702
22.9581 | -35.5571 | -33.8762
22.9642 | -35.5577 | -33.8828
22.9703 | -35.5582 | -33.8889
22.9764 | -35.5587 | -33.8954
22.9825 | -35.5592 | -33.9013
22.9886 | -35.5597 | -33.9076
22.9948 | -35.5602 | -33.9138
23.0009 | -35.5607 | -33.9199
23.007 | -35.5613 | -33.9261
23.0131 | -35.5618 | -33.9324
23.0192 | -35.5623 | -33.9383
23.0253 | -35.5629 | -33.9447
23.0314 | -35.5633 | -33.951
23.0375 | -35.5639 | -33.9572
23.0436 | -35.5643 | -33.9629
23.0497 | -35.5648 | -33.9693
23.0558 | -35.5653 | -33.9752
23.0619 | -35.5658 | -33.9814
23.068 | -35.5663 | -33.9877
23.0741 | -35.5668 | -33.9933
23.0802 | -35.5673 | -33.9996
23.0863 | -35.5678 | -34.0058
23.0924 | -35.5683 | -34.0121
23.0985 | -35.5688 | -34.0183
23.1046 | -35.5692 | -34.0245
23.1107 | -35.5698 | -34.0302
23.1168 | -35.5702 | -34.0364
23.1229 | -35.5707 | -34.0423
23.129 | -35.5712 | -34.0479
23.1351 | -35.5718 | -34.054
23.1413 | -35.5722 | -34.0603
23.1474 | -35.5727 | -34.0663
23.1535 | -35.5731 | -34.0723
23.1596 | -35.5736 | -34.0784
23.1657 | -35.5741 | -34.0848
23.1718 | -35.5746 | -34.091
23.1779 | -35.575 | -34.0975
23.184 | -35.5755 | -34.1035
23.1901 | -35.576 | -34.1099
23.1962 | -35.5766 | -34.1162
23.2023 | -35.577 | -34.1223
23.2084 | -35.5775 | -34.1287
23.2145 | -35.578 | -34.1345
23.2206 | -35.5785 | -34.1407
23.2267 | -35.5791 | -34.1468
23.2328 | -35.5795 | -34.153
23.2389 | -35.58 | -34.1588
23.245 | -35.5805 | -34.1652
23.2511 | -35.581 | -34.1708
23.2572 | -35.5815 | -34.1771
23.2633 | -35.582 | -34.183
23.2694 | -35.5825 | -34.1894
23.2755 | -35.583 | -34.1951
23.2817 | -35.5835 | -34.2008
23.2878 | -35.584 | -34.2071
23.2939 | -35.5845 | -34.2133
23.3 | -35.585 | -34.2192
23.3061 | -35.5855 | -34.2252
23.3122 | -35.586 | -34.2313
23.3183 | -35.5864 | -34.2372
23.3244 | -35.5869 | -34.2431
23.3305 | -35.5875 | -34.2492
23.3366 | -35.588 | -34.2557
23.3427 | -35.5885 | -34.2618
23.3488 | -35.589 | -34.2677
23.3549 | -35.5895 | -34.2736
23.361 | -35.59 | -34.2796
23.3671 | -35.5905 | -34.286
23.3732 | -35.591 | -34.292
23.3793 | -35.5915 | -34.2981
23.3854 | -35.592 | -34.3045
23.3915 | -35.5925 | -34.3102
23.3976 | -35.593 | -34.3165
23.4037 | -35.5935 | -34.3226
23.4098 | -35.5941 | -34.3292
23.4159 | -35.5945 | -34.3351
23.422 | -35.595 | -34.3414
23.4282 | -35.5955 | -34.347
23.4343 | -35.5961 | -34.3529
23.4404 | -35.5966 | -34.359
23.4465 | -35.5971 | -34.3645
23.4526 | -35.5976 | -34.3707
23.4587 | -35.598 | -34.377
23.4648 | -35.5985 | -34.3828
23.4709 | -35.599 | -34.3888
23.477 | -35.5995 | -34.3948
23.4831 | -35.5999 | -34.4007
23.4892 | -35.6004 | -34.4065
23.4953 | -35.6009 | -34.4126
23.5014 | -35.6014 | -34.4182
23.5075 | -35.6018 | -34.4241
23.5136 | -35.6023 | -34.4305
23.5197 | -35.6028 | -34.4361
23.5258 | -35.6033 | -34.4418
23.5319 | -35.6037 | -34.4476
23.538 | -35.6042 | -34.4534
23.5441 | -35.6047 | -34.4596
23.5502 | -35.6052 | -34.4655
23.5563 | -35.6057 | -34.4719
23.5624 | -35.6061 | -34.4779
23.5686 | -35.6066 | -34.4829
23.5747 | -35.6071 | -34.489
23.5808 | -35.6076 | -34.495
23.5869 | -35.6081 | -34.5007
23.593 | -35.6086 | -34.5066
23.5991 | -35.6091 | -34.5123
23.6052 | -35.6095 | -34.5176
23.6113 | -35.61 | -34.5236
23.6174 | -35.6104 | -34.5294
23.6235 | -35.6109 | -34.5356
23.6296 | -35.6114 | -34.5416
23.6357 | -35.6119 | -34.548
23.6418 | -35.6123 | -34.5533
23.6479 | -35.6128 | -34.5593
23.654 | -35.6133 | -34.5651
23.6601 | -35.6137 | -34.5711
23.6662 | -35.6142 | -34.5774
23.6723 | -35.6146 | -34.5832
23.6784 | -35.6151 | -34.5891
23.6845 | -35.6155 | -34.5953
23.6906 | -35.616 | -34.6015
23.6967 | -35.6165 | -34.6075
23.7028 | -35.617 | -34.6132
23.7089 | -35.6175 | -34.6187
23.7151 | -35.618 | -34.6247
23.7212 | -35.6184 | -34.6307
23.7273 | -35.6189 | -34.6367
23.7334 | -35.6194 | -34.6428
23.7395 | -35.6199 | -34.6488
23.7456 | -35.6203 | -34.6537
23.7517 | -35.6208 | -34.6596
23.7578 | -35.6213 | -34.6654
23.7639 | -35.6217 | -34.6711
23.77 | -35.6222 | -34.6774
23.7761 | -35.6227 | -34.6831
23.7822 | -35.6231 | -34.6884
23.7883 | -35.6236 | -34.6942
23.7944 | -35.6241 | -34.6999
23.8005 | -35.6246 | -34.7057
23.8066 | -35.625 | -34.7117
23.8127 | -35.6255 | -34.7174
23.8188 | -35.626 | -34.7236
23.8249 | -35.6265 | -34.73
23.831 | -35.6269 | -34.736
23.8371 | -35.6274 | -34.7418
23.8432 | -35.6279 | -34.7477
23.8493 | -35.6284 | -34.7532
23.8555 | -35.6288 | -34.7592
23.8616 | -35.6293 | -34.7651
23.8677 | -35.6297 | -34.7704
23.8738 | -35.6302 | -34.7764
23.8799 | -35.6307 | -34.7821
23.886 | -35.631 | -34.7877
23.8921 | -35.6315 | -34.7936
23.8982 | -35.632 | -34.7997
23.9043 | -35.6325 | -34.8059
23.9104 | -35.633 | -34.8123
23.9165 | -35.6334 | -34.8179
23.9226 | -35.6339 | -34.8226
23.9287 | -35.6343 | -34.8287
23.9348 | -35.6348 | -34.8348
23.9409 | -35.6353 | -34.8404
23.947 | -35.6357 | -34.8465
23.9531 | -35.6361 | -34.8516
23.9592 | -35.6365 | -34.8572
23.9653 | -35.637 | -34.863
23.9714 | -35.6375 | -34.8686
23.9775 | -35.638 | -34.8744
23.9836 | -35.6384 | -34.8801
23.9897 | -35.6389 | -34.8854
23.9958 | -35.6393 | -34.8916
24.002 | -35.6398 | -34.8975
24.0081 | -35.6403 | -34.9038
24.0142 | -35.6407 | -34.9091
24.0203 | -35.6412 | -34.9148
24.0264 | -35.6416 | -34.9203
24.0325 | -35.6421 | -34.9259
24.0386 | -35.6427 | -34.9321
24.0447 | -35.6431 | -34.9374
24.0508 | -35.6436 | -34.9433
24.0569 | -35.6441 | -34.9491
24.063 | -35.6445 | -34.9541
24.0691 | -35.645 | -34.9596
24.0752 | -35.6455 | -34.9652
24.0813 | -35.6459 | -34.9709
24.0874 | -35.6464 | -34.9768
24.0935 | -35.6469 | -34.9828
24.0996 | -35.6473 | -34.9879
24.1057 | -35.6478 | -34.9937
24.1118 | -35.6483 | -34.9999
24.1179 | -35.6488 | -35.0056
24.124 | -35.6493 | -35.0114
24.1301 | -35.6497 | -35.0172
24.1362 | -35.6501 | -35.023
24.1424 | -35.6507 | -35.0291
24.1485 | -35.6512 | -35.0352
24.1546 | -35.6516 | -35.0407
24.1607 | -35.6521 | -35.0464
24.1668 | -35.6526 | -35.052
24.1729 | -35.653 | -35.0575
24.179 | -35.6535 | -35.0632
24.1851 | -35.6539 | -35.0688
24.1912 | -35.6544 | -35.0749
24.1973 | -35.6549 | -35.0805
24.2034 | -35.6553 | -35.0861
24.2095 | -35.6557 | -35.0919
24.2156 | -35.6562 | -35.0978
24.2217 | -35.6566 | -35.1034
24.2278 | -35.657 | -35.1091
24.2339 | -35.6575 | -35.1145
24.24 | -35.6579 | -35.1201
24.2461 | -35.6584 | -35.1257
24.2522 | -35.6588 | -35.1314
24.2583 | -35.6593 | -35.1375
24.2644 | -35.6597 | -35.1434
24.2705 | -35.6602 | -35.1488
24.2766 | -35.6606 | -35.1544
24.2827 | -35.6611 | -35.1603
24.2889 | -35.6615 | -35.1662
24.295 | -35.662 | -35.1718
24.3011 | -35.6624 | -35.1776
24.3072 | -35.6628 | -35.1835
24.3133 | -35.6632 | -35.1892
24.3194 | -35.6636 | -35.1947
24.3255 | -35.664 | -35.2006
24.3316 | -35.6645 | -35.2064
24.3377 | -35.665 | -35.2124
24.3438 | -35.6653 | -35.2178
24.3499 | -35.6658 | -35.2237
24.356 | -35.6663 | -35.2293
24.3621 | -35.6668 | -35.2348
24.3682 | -35.6672 | -35.2403
24.3743 | -35.6677 | -35.2459
24.3804 | -35.6681 | -35.2517
24.3865 | -35.6685 | -35.2572
24.3926 | -35.6689 | -35.263
24.3987 | -35.6694 | -35.2682
24.4048 | -35.6698 | -35.274
24.4109 | -35.6702 | -35.2795
24.417 | -35.6707 | -35.2849
24.4231 | -35.6711 | -35.2905
24.4293 | -35.6716 | -35.2964
24.4354 | -35.6721 | -35.3019
24.4415 | -35.6725 | -35.3074
24.4476 | -35.6728 | -35.3124
24.4537 | -35.6733 | -35.3178
24.4598 | -35.6738 | -35.323
24.4659 | -35.6742 | -35.3282
24.472 | -35.6747 | -35.3334
24.4781 | -35.6751 | -35.3393
24.4842 | -35.6755 | -35.3445
24.4903 | -35.6759 | -35.35
24.4964 | -35.6764 | -35.3555
24.5025 | -35.6769 | -35.3611
24.5086 | -35.6774 | -35.367
24.5147 | -35.6778 | -35.3724
24.5208 | -35.6782 | -35.3774
24.5269 | -35.6786 | -35.3831
24.533 | -35.6791 | -35.389
24.5391 | -35.6795 | -35.3945
24.5452 | -35.6799 | -35.4001
24.5513 | -35.6803 | -35.4056
24.5574 | -35.6808 | -35.4117
24.5635 | -35.6812 | -35.4171
24.5696 | -35.6816 | -35.4225
24.5758 | -35.6821 | -35.4282
24.5819 | -35.6825 | -35.4337
24.588 | -35.683 | -35.4389
24.5941 | -35.6835 | -35.4441
24.6002 | -35.6839 | -35.4497
24.6063 | -35.6843 | -35.4551
24.6124 | -35.6848 | -35.4604
24.6185 | -35.6853 | -35.4663
24.6246 | -35.6857 | -35.4713
24.6307 | -35.6861 | -35.477
24.6368 | -35.6865 | -35.4826
24.6429 | -35.687 | -35.4886
24.649 | -35.6875 | -35.4941
24.6551 | -35.6879 | -35.4997
24.6612 | -35.6883 | -35.5048
24.6673 | -35.6888 | -35.5103
24.6734 | -35.6892 | -35.5161
24.6795 | -35.6896 | -35.5219
24.6856 | -35.6901 | -35.5276
24.6917 | -35.6905 | -35.5333
24.6978 | -35.6908 | -35.5383
24.7039 | -35.6913 | -35.5438
24.71 | -35.6917 | -35.5494
24.7162 | -35.6921 | -35.5547
24.7223 | -35.6926 | -35.5603
24.7284 | -35.693 | -35.5656
24.7345 | -35.6934 | -35.5712
24.7406 | -35.6939 | -35.5762
24.7467 | -35.6943 | -35.5818
24.7528 | -35.6947 | -35.5871
24.7589 | -35.6952 | -35.5925
24.765 | -35.6955 | -35.5979
24.7711 | -35.6959 | -35.6029
24.7772 | -35.6964 | -35.6084
24.7833 | -35.6969 | -35.6142
24.7894 | -35.6972 | -35.6197
24.7955 | -35.6977 | -35.6252
24.8016 | -35.6981 | -35.63
24.8077 | -35.6985 | -35.6355
24.8138 | -35.699 | -35.6419
24.8199 | -35.6994 | -35.6469
24.826 | -35.6998 | -35.6524
24.8321 | -35.7003 | -35.6577
24.8382 | -35.7006 | -35.6629
24.8443 | -35.7011 | -35.6677
24.8504 | -35.7016 | -35.6732
24.8565 | -35.702 | -35.6791
24.8627 | -35.7025 | -35.6842
24.8688 | -35.703 | -35.6899
24.8749 | -35.7034 | -35.6953
24.881 | -35.7038 | -35.7005
24.8871 | -35.7043 | -35.7058
24.8932 | -35.7048 | -35.7113
24.8993 | -35.7052 | -35.7164
24.9054 | -35.7057 | -35.7214
24.9115 | -35.7061 | -35.7268
24.9176 | -35.7065 | -35.7324
24.9237 | -35.7069 | -35.7382
24.9298 | -35.7073 | -35.7436
24.9359 | -35.7077 | -35.7491
24.942 | -35.708 | -35.7544
24.9481 | -35.7084 | -35.7598
24.9542 | -35.7088 | -35.7652
24.9603 | -35.7092 | -35.7706
24.9664 | -35.7097 | -35.7758
24.9725 | -35.7101 | -35.7811
24.9786 | -35.7105 | -35.7856
24.9847 | -35.711 | -35.7912
24.9908 | -35.7114 | -35.797
24.9969 | -35.7118 | -35.802
25.0031 | -35.7122 | -35.8071
25.0092 | -35.7127 | -35.8123
25.0153 | -35.7131 | -35.8175
25.0214 | -35.7136 | -35.8227
25.0275 | -35.7141 | -35.8281
25.0336 | -35.7146 | -35.8339
25.0397 | -35.715 | -35.8393
25.0458 | -35.7156 | -35.8447
25.0519 | -35.7159 | -35.85
25.058 | -35.7163 | -35.8554
25.0641 | -35.7167 | -35.8604
25.0702 | -35.717 | -35.8655
25.0763 | -35.7175 | -35.8715
25.0824 | -35.7179 | -35.8768
25.0885 | -35.7183 | -35.8821
25.0946 | -35.7188 | -35.8874
25.1007 | -35.7192 | -35.8926
25.1068 | -35.7196 | -35.8977
25.1129 | -35.72 | -35.9029
25.119 | -35.7204 | -35.9078
25.1251 | -35.7208 | -35.9131
25.1312 | -35.7212 | -35.9191
25.1373 | -35.7216 | -35.9243
25.1435 | -35.722 | -35.9302
25.1496 | -35.7224 | -35.9346
25.1557 | -35.7228 | -35.9399
25.1618 | -35.7232 | -35.9456
25.1679 | -35.7236 | -35.9513
25.174 | -35.7239 | -35.9567
25.1801 | -35.7243 | -35.9621
25.1862 | -35.7247 | -35.9675
25.1923 | -35.725 | -35.9726
25.1984 | -35.7254 | -35.9779
25.2045 | -35.7257 | -35.9832
25.2106 | -35.7261 | -35.989
25.2167 | -35.7265 | -35.9945
25.2228 | -35.7269 | -35.9992
25.2289 | -35.7274 | -36.0046
25.235 | -35.7278 | -36.0101
25.2411 | -35.7282 | -36.016
25.2472 | -35.7286 | -36.0213
25.2533 | -35.729 | -36.0269
25.2594 | -35.7293 | -36.0317
25.2655 | -35.7298 | -36.0369
25.2716 | -35.7302 | -36.0422
25.2777 | -35.7306 | -36.0475
25.2838 | -35.731 | -36.0533
25.29 | -35.7314 | -36.0585
25.2961 | -35.7318 | -36.0635
25.3022 | -35.7323 | -36.0687
25.3083 | -35.7326 | -36.0741
25.3144 | -35.733 | -36.0792
25.3205 | -35.7334 | -36.0845
25.3266 | -35.7338 | -36.0897
25.3327 | -35.7342 | -36.0947
25.3388 | -35.7346 | -36.1001
25.3449 | -35.735 | -36.1054
25.351 | -35.7354 | -36.1108
25.3571 | -35.7359 | -36.1166
25.3632 | -35.7362 | -36.1218
25.3693 | -35.7366 | -36.1272
25.3754 | -35.737 | -36.1325
25.3815 | -35.7374 | -36.1375
25.3876 | -35.7378 | -36.1427
25.3937 | -35.7381 | -36.1478
25.3998 | -35.7384 | -36.1523
25.4059 | -35.7388 | -36.1575
25.412 | -35.7392 | -36.1623
25.4181 | -35.7396 | -36.167
25.4242 | -35.74 | -36.1721
25.4304 | -35.7404 | -36.1773
25.4365 | -35.7408 | -36.1823
25.4426 | -35.7411 | -36.1875
25.4487 | -35.7414 | -36.193
25.4548 | -35.7418 | -36.1982
25.4609 | -35.7423 | -36.2035
25.467 | -35.7426 | -36.2087
25.4731 | -35.743 | -36.2142
25.4792 | -35.7434 | -36.2195
25.4853 | -35.7438 | -36.2247
25.4914 | -35.7442 | -36.2303
25.4975 | -35.7446 | -36.2358
25.5036 | -35.745 | -36.2406
25.5097 | -35.7454 | -36.2459
25.5158 | -35.7458 | -36.251
25.5219 | -35.7462 | -36.256
25.528 | -35.7466 | -36.2611
25.5341 | -35.747 | -36.2667
25.5402 | -35.7474 | -36.2723
25.5463 | -35.7478 | -36.2776
25.5524 | -35.7482 | -36.2828
25.5585 | -35.7486 | -36.2878
25.5646 | -35.749 | -36.2931
25.5707 | -35.7494 | -36.298
25.5769 | -35.7496 | -36.3029
25.583 | -35.7501 | -36.3084
25.5891 | -35.7504 | -36.3134
25.5952 | -35.7508 | -36.3186
25.6013 | -35.7512 | -36.3236
25.6074 | -35.7516 | -36.3283
25.6135 | -35.7519 | -36.3337
25.6196 | -35.7524 | -36.3392
25.6257 | -35.7528 | -36.3444
25.6318 | -35.7532 | -36.3493
25.6379 | -35.7536 | -36.3544
25.644 | -35.7539 | -36.3595
25.6501 | -35.7543 | -36.3651
25.6562 | -35.7547 | -36.3704
25.6623 | -35.7551 | -36.3753
25.6684 | -35.7555 | -36.3806
25.6745 | -35.7559 | -36.3857
25.6806 | -35.7563 | -36.3901
25.6867 | -35.7567 | -36.3953
25.6928 | -35.757 | -36.4007
25.6989 | -35.7573 | -36.4053
25.705 | -35.7578 | -36.41
25.7111 | -35.7581 | -36.4153
25.7173 | -35.7585 | -36.4196
25.7234 | -35.7589 | -36.4248
25.7295 | -35.7592 | -36.43
25.7356 | -35.7596 | -36.4351
25.7417 | -35.76 | -36.4403
25.7478 | -35.7604 | -36.4455
25.7539 | -35.7608 | -36.4505
25.76 | -35.7611 | -36.4556
25.7661 | -35.7615 | -36.4609
25.7722 | -35.7619 | -36.4662
25.7783 | -35.7623 | -36.4713
25.7844 | -35.7627 | -36.4766
25.7905 | -35.7631 | -36.4818
25.7966 | -35.7635 | -36.4871
25.8027 | -35.7638 | -36.492
25.8088 | -35.7642 | -36.4975
25.8149 | -35.7645 | -36.5029
25.821 | -35.7649 | -36.5076
25.8271 | -35.7653 | -36.5123
25.8332 | -35.7656 | -36.5176
25.8393 | -35.766 | -36.5229
25.8454 | -35.7664 | -36.5278
25.8515 | -35.7667 | -36.5331
25.8576 | -35.7671 | -36.5379
25.8638 | -35.7675 | -36.5428
25.8699 | -35.7678 | -36.548
25.876 | -35.7682 | -36.5532
25.8821 | -35.7686 | -36.5582
25.8882 | -35.7689 | -36.5632
25.8943 | -35.7693 | -36.5688
25.9004 | -35.7697 | -36.5736
25.9065 | -35.7701 | -36.5783
25.9126 | -35.7705 | -36.5832
25.9187 | -35.7708 | -36.5883
25.9248 | -35.7711 | -36.5934
25.9309 | -35.7715 | -36.5987
25.937 | -35.7719 | -36.6039
25.9431 | -35.7723 | -36.609
25.9492 | -35.7726 | -36.6144
25.9553 | -35.773 | -36.6196
25.9614 | -35.7735 | -36.6244
25.9675 | -35.7738 | -36.6296
25.9736 | -35.7742 | -36.635
25.9797 | -35.7746 | -36.64
25.9858 | -35.7749 | -36.6452
25.9919 | -35.7753 | -36.6505
25.998 | -35.7756 | -36.6553
26.0042 | -35.776 | -36.6604
26.0103 | -35.7764 | -36.6658
26.0164 | -35.7768 | -36.6708
26.0225 | -35.7771 | -36.6761
26.0286 | -35.7775 | -36.6814
26.0347 | -35.7779 | -36.6861
26.0408 | -35.7783 | -36.691
26.0469 | -35.7786 | -36.6964
26.053 | -35.779 | -36.7014
26.0591 | -35.7794 | -36.7068
26.0652 | -35.7797 | -36.7117
26.0713 | -35.7801 | -36.7165
26.0774 | -35.7804 | -36.7217
26.0835 | -35.7808 | -36.7269
26.0896 | -35.7812 | -36.7318
26.0957 | -35.7815 | -36.7366
26.1018 | -35.7818 | -36.7415
26.1079 | -35.7822 | -36.7466
26.114 | -35.7826 | -36.7517
26.1201 | -35.7829 | -36.7567
26.1262 | -35.7833 | -36.7614
26.1323 | -35.7836 | -36.7663
26.1384 | -35.784 | -36.7709
26.1445 | -35.7843 | -36.7757
26.1507 | -35.7847 | -36.7809
26.1568 | -35.7851 | -36.7862
26.1629 | -35.7854 | -36.791
26.169 | -35.7858 | -36.796
26.1751 | -35.7862 | -36.8008
26.1812 | -35.7865 | -36.8056
26.1873 | -35.7869 | -36.811
26.1934 | -35.7873 | -36.8159
26.1995 | -35.7877 | -36.821
26.2056 | -35.788 | -36.8256
26.2117 | -35.7884 | -36.8305
26.2178 | -35.7887 | -36.8356
26.2239 | -35.7891 | -36.8407
26.23 | -35.7895 | -36.8458
26.2361 | -35.7898 | -36.8511
26.2422 | -35.7902 | -36.8557
26.2483 | -35.7906 | -36.8608
26.2544 | -35.7909 | -36.8656
26.2605 | -35.7913 | -36.871
26.2666 | -35.7917 | -36.8758
26.2727 | -35.7921 | -36.8811
26.2788 | -35.7925 | -36.8858
26.2849 | -35.7928 | -36.891
26.2911 | -35.7932 | -36.8962
26.2972 | -35.7936 | -36.9009
26.3033 | -35.794 | -36.9058
26.3094 | -35.7943 | -36.9106
26.3155 | -35.7947 | -36.9157
26.3216 | -35.795 | -36.9211
26.3277 | -35.7954 | -36.926
26.3338 | -35.7958 | -36.931
26.3399 | -35.7962 | -36.9358
26.346 | -35.7965 | -36.9409
26.3521 | -35.7968 | -36.9461
26.3582 | -35.7972 | -36.9512
26.3643 | -35.7975 | -36.9562
26.3704 | -35.7979 | -36.9612
26.3765 | -35.7984 | -36.9662
26.3826 | -35.7986 | -36.9706
26.3887 | -35.799 | -36.976
26.3948 | -35.7993 | -36.981
26.4009 | -35.7997 | -36.9863
26.407 | -35.8001 | -36.9913
26.4131 | -35.8004 | -36.9964
26.4192 | -35.8008 | -37.0015
26.4253 | -35.8011 | -37.0069
26.4314 | -35.8015 | -37.012
26.4376 | -35.8018 | -37.0168
26.4437 | -35.8021 | -37.0224
26.4498 | -35.8025 | -37.0275
26.4559 | -35.8028 | -37.0318
26.462 | -35.8032 | -37.037
26.4681 | -35.8035 | -37.0421
26.4742 | -35.8038 | -37.0473
26.4803 | -35.8042 | -37.0522
26.4864 | -35.8045 | -37.0574
26.4925 | -35.8049 | -37.0624
26.4986 | -35.8053 | -37.0672
26.5047 | -35.8056 | -37.0723
26.5108 | -35.8059 | -37.0771
26.5169 | -35.8062 | -37.082
26.523 | -35.8066 | -37.0867
26.5291 | -35.807 | -37.0916
26.5352 | -35.8074 | -37.0963
26.5413 | -35.8077 | -37.1013
26.5474 | -35.8081 | -37.1065
26.5535 | -35.8085 | -37.1117
26.5596 | -35.8088 | -37.1167
26.5657 | -35.8091 | -37.1217
26.5718 | -35.8095 | -37.1267
26.578 | -35.8099 | -37.1319
26.5841 | -35.8102 | -37.1367
26.5902 | -35.8105 | -37.1416
26.5963 | -35.8108 | -37.1463
26.6024 | -35.8112 | -37.1515
26.6085 | -35.8115 | -37.1565
26.6146 | -35.8119 | -37.1617
26.6207 | -35.8123 | -37.1666
26.6268 | -35.8126 | -37.1716
26.6329 | -35.8129 | -37.1762
26.639 | -35.8133 | -37.1812
26.6451 | -35.8136 | -37.1865
26.6512 | -35.814 | -37.1916
26.6573 | -35.8143 | -37.1968
26.6634 | -35.8147 | -37.2015
26.6695 | -35.815 | -37.2069
26.6756 | -35.8154 | -37.212
26.6817 | -35.8158 | -37.2168
26.6878 | -35.8161 | -37.222
26.6939 | -35.8165 | -37.2267
26.7 | -35.8168 | -37.2315
26.7061 | -35.8171 | -37.2362
26.7122 | -35.8175 | -37.2412
26.7183 | -35.8179 | -37.2462
26.7245 | -35.8182 | -37.2512
26.7306 | -35.8185 | -37.2564
26.7367 | -35.8189 | -37.2608
26.7428 | -35.8192 | -37.266
26.7489 | -35.8196 | -37.271
26.755 | -35.8199 | -37.2761
26.7611 | -35.8203 | -37.2811
26.7672 | -35.8206 | -37.286
26.7733 | -35.8209 | -37.2908
26.7794 | -35.8213 | -37.2961
26.7855 | -35.8216 | -37.3014
26.7916 | -35.822 | -37.3064
26.7977 | -35.8224 | -37.3115
26.8038 | -35.8227 | -37.3163
26.8099 | -35.823 | -37.3216
26.816 | -35.8234 | -37.3267
26.8221 | -35.8237 | -37.3316
26.8282 | -35.8241 | -37.3369
26.8343 | -35.8245 | -37.3416
26.8404 | -35.8247 | -37.3463
26.8465 | -35.8251 | -37.3515
26.8526 | -35.8254 | -37.3566
26.8587 | -35.8257 | -37.3618
26.8649 | -35.826 | -37.3669
26.871 | -35.8263 | -37.3717
26.8771 | -35.8266 | -37.3767
26.8832 | -35.827 | -37.3813
26.8893 | -35.8273 | -37.3866
26.8954 | -35.8276 | -37.3914
26.9015 | -35.8279 | -37.3961
26.9076 | -35.8282 | -37.4011
26.9137 | -35.8285 | -37.4058
26.9198 | -35.8289 | -37.4108
26.9259 | -35.8292 | -37.4159
26.932 | -35.8296 | -37.4208
26.9381 | -35.8299 | -37.426
26.9442 | -35.8303 | -37.4307
26.9503 | -35.8306 | -37.4358
26.9564 | -35.831 | -37.4408
26.9625 | -35.8313 | -37.4458
26.9686 | -35.8316 | -37.4506
26.9747 | -35.8319 | -37.456
26.9808 | -35.8323 | -37.4613
26.9869 | -35.8326 | -37.4663
26.993 | -35.8329 | -37.4713
26.9991 | -35.8333 | -37.4762
27.0052 | -35.8337 | -37.4813
27.0114 | -35.834 | -37.4866
27.0175 | -35.8343 | -37.4917
27.0236 | -35.8346 | -37.4966
27.0297 | -35.8349 | -37.5014
27.0358 | -35.8353 | -37.5064
27.0419 | -35.8356 | -37.5116
27.048 | -35.836 | -37.5164
27.0541 | -35.8363 | -37.5209
27.0602 | -35.8366 | -37.5258
27.0663 | -35.8369 | -37.5311
27.0724 | -35.8373 | -37.5361
27.0785 | -35.8376 | -37.5414
27.0846 | -35.8381 | -37.5462
27.0907 | -35.8384 | -37.5513
27.0968 | -35.8387 | -37.5566
27.1029 | -35.8391 | -37.5613
27.109 | -35.8394 | -37.5661
27.1151 | -35.8397 | -37.5712
27.1212 | -35.84 | -37.5759
27.1273 | -35.8404 | -37.5807
27.1334 | -35.8407 | -37.5855
27.1395 | -35.8411 | -37.5907
27.1456 | -35.8415 | -37.5956
27.1518 | -35.8418 | -37.601
27.1579 | -35.8421 | -37.6057
27.164 | -35.8425 | -37.6106
27.1701 | -35.8429 | -37.6156
27.1762 | -35.8432 | -37.6204
27.1823 | -35.8435 | -37.6254
27.1884 | -35.8439 | -37.6305
27.1945 | -35.8442 | -37.6351
27.2006 | -35.8445 | -37.6401
27.2067 | -35.8448 | -37.6449
27.2128 | -35.8451 | -37.6499
27.2189 | -35.8454 | -37.655
27.225 | -35.8458 | -37.6601
27.2311 | -35.8461 | -37.665
27.2372 | -35.8464 | -37.6698
27.2433 | -35.8468 | -37.6747
27.2494 | -35.8471 | -37.6797
27.2555 | -35.8475 | -37.6848
27.2616 | -35.8478 | -37.6894
27.2677 | -35.8481 | -37.6946
27.2738 | -35.8484 | -37.6995
27.2799 | -35.8488 | -37.7044
27.286 | -35.8491 | -37.7094
27.2921 | -35.8494 | -37.7143
27.2983 | -35.8497 | -37.719
27.3044 | -35.8501 | -37.7238
27.3105 | -35.8504 | -37.7288
27.3166 | -35.8508 | -37.734
27.3227 | -35.8511 | -37.7388
27.3288 | -35.8515 | -37.7441
27.3349 | -35.8518 | -37.7493
27.341 | -35.8521 | -37.7542
27.3471 | -35.8525 | -37.759
27.3532 | -35.8528 | -37.7637
27.3593 | -35.8532 | -37.7688
27.3654 | -35.8535 | -37.7738
27.3715 | -35.8538 | -37.7789
27.3776 | -35.8541 | -37.784
27.3837 | -35.8544 | -37.7887
27.3898 | -35.8547 | -37.7937
27.3959 | -35.855 | -37.7986
27.402 | -35.8554 | -37.8037
27.4081 | -35.8557 | -37.8086
27.4142 | -35.8561 | -37.8136
27.4203 | -35.8564 | -37.8182
27.4264 | -35.8567 | -37.8234
27.4325 | -35.857 | -37.8285
27.4387 | -35.8573 | -37.8328
27.4448 | -35.8576 | -37.8375
27.4509 | -35.858 | -37.8425
27.457 | -35.8583 | -37.8473
27.4631 | -35.8587 | -37.8526
27.4692 | -35.859 | -37.8576
27.4753 | -35.8593 | -37.8626
27.4814 | -35.8596 | -37.8675
27.4875 | -35.86 | -37.8729
27.4936 | -35.8603 | -37.8775
27.4997 | -35.8606 | -37.8823
27.5058 | -35.861 | -37.8874
27.5119 | -35.8613 | -37.8921
27.518 | -35.8616 | -37.8969
27.5241 | -35.862 | -37.9019
27.5302 | -35.8623 | -37.9072
27.5363 | -35.8627 | -37.9123
27.5424 | -35.863 | -37.9173
27.5485 | -35.8633 | -37.9223
27.5546 | -35.8636 | -37.9271
27.5607 | -35.8639 | -37.9321
27.5668 | -35.8642 | -37.937
27.5729 | -35.8645 | -37.9421
27.5791 | -35.8648 | -37.9476
27.5852 | -35.8651 | -37.9526
27.5913 | -35.8654 | -37.9573
27.5974 | -35.8658 | -37.9624
27.6035 | -35.8661 | -37.9675
27.6096 | -35.8664 | -37.9725
27.6157 | -35.8667 | -37.9771
27.6218 | -35.867 | -37.9824
27.6279 | -35.8674 | -37.9872
27.634 | -35.8677 | -37.992
27.6401 | -35.868 | -37.997
27.6462 | -35.8684 | -38.0021
27.6523 | -35.8687 | -38.007
27.6584 | -35.869 | -38.012
27.6645 | -35.8693 | -38.0166
27.6706 | -35.8696 | -38.022
27.6767 | -35.87 | -38.027
27.6828 | -35.8703 | -38.0318
27.6889 | -35.8706 | -38.0369
27.695 | -35.8709 | -38.0418
27.7011 | -35.8712 | -38.0468
27.7072 | -35.8715 | -38.0518
27.7133 | -35.8719 | -38.0569
27.7194 | -35.8722 | -38.0622
27.7256 | -35.8725 | -38.0672
27.7317 | -35.8728 | -38.0722
27.7378 | -35.8732 | -38.0771
27.7439 | -35.8735 | -38.0822
27.75 | -35.8739 | -38.0873
27.7561 | -35.8742 | -38.0922
27.7622 | -35.8746 | -38.0975
27.7683 | -35.8749 | -38.1025
27.7744 | -35.8753 | -38.1073
27.7805 | -35.8756 | -38.112
27.7866 | -35.876 | -38.1172
27.7927 | -35.8762 | -38.122
27.7988 | -35.8765 | -38.1268
27.8049 | -35.8769 | -38.1316
27.811 | -35.8772 | -38.1366
27.8171 | -35.8775 | -38.1416
27.8232 | -35.8778 | -38.1465
27.8293 | -35.8781 | -38.1514
27.8354 | -35.8784 | -38.1564
27.8415 | -35.8787 | -38.1613
27.8476 | -35.879 | -38.1666
27.8537 | -35.8794 | -38.1713
27.8598 | -35.8797 | -38.1761
27.866 | -35.88 | -38.1811
27.8721 | -35.8803 | -38.1857
27.8782 | -35.8806 | -38.1906
27.8843 | -35.8809 | -38.1953
27.8904 | -35.8813 | -38.2002
27.8965 | -35.8817 | -38.2053
27.9026 | -35.8819 | -38.2104
27.9087 | -35.8822 | -38.2153
27.9148 | -35.8825 | -38.2201
27.9209 | -35.8829 | -38.2251
27.927 | -35.8832 | -38.2303
27.9331 | -35.8835 | -38.2354
27.9392 | -35.8839 | -38.2403
27.9453 | -35.8841 | -38.2453
27.9514 | -35.8845 | -38.2504
27.9575 | -35.8848 | -38.2556
27.9636 | -35.8851 | -38.2606
27.9697 | -35.8855 | -38.2656
27.9758 | -35.8858 | -38.2707
27.9819 | -35.8861 | -38.2757
27.988 | -35.8864 | -38.2807
27.9941 | -35.8868 | -38.2856
28.0002 | -35.887 | -38.2904
28.0063 | -35.8873 | -38.2955
28.0125 | -35.8876 | -38.3005
28.0186 | -35.8879 | -38.3054
28.0247 | -35.8883 | -38.3104
28.0308 | -35.8886 | -38.3151
28.0369 | -35.8889 | -38.3202
28.043 | -35.8892 | -38.3246
28.0491 | -35.8896 | -38.3295
28.0552 | -35.8899 | -38.3344
28.0613 | -35.8902 | -38.3393
28.0674 | -35.8905 | -38.3442
28.0735 | -35.8908 | -38.3493

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.6384 | -42.8489
21.4908 | -35.6595 | -35.5912
16.4948 | -35.018 | -28.1456
12.6602 | -34.6402 | -21.8086
9.7171 | -34.4134 | -16.8593
7.4582 | -34.2516 | -12.8357
5.7244 | -34.1311 | -9.5795
4.3936 | -34.0504 | -6.9524
3.3722 | -34.0335 | -4.8779
2.5883 | -34.0088 | -3.3058
1.9866 | -34.0035 | -2.0512
1.5248 | -33.9959 | -1.0442
1.1703 | -34.0054 | -0.38827
1 | -34.0022 | -0.062181
0.89822 | -33.991 | 0.21319
0.68941 | -33.9688 | 0.79948
0.52919 | -33.9684 | 1.2523
0.40612 | -33.9743 | 1.8671
0.31171 | -33.9601 | 2.3212
0.23928 | -33.9659 | 3.0481
0.18358 | -33.9686 | 3.4779
0.14091 | -33.9685 | 4.37
0.10819 | -34.0086 | 5.5211
0.083017 | -34.0101 | 7.0135
0.063705 | -34.0506 | 8.8883
0.048923 | -34.1075 | 11.3947
0.037479 | -34.2127 | 14.5973
0.028801 | -34.402 | 18.62
0.022078 | -34.7063 | 23.9478
0.017023 | -35.1765 | 30.3308
0.013018 | -35.9868 | 38.0009
0.010014 | -37.0969 | 45.8926

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003704 | -8.1415
0.007408 | -25.8098
0.011112 | -34.9514
0.014816 | -40.7768
0.018519 | -43.4558
0.022223 | -45.9248
0.025927 | -48.09
0.029631 | -49.3516
0.033335 | -50.2078
0.037039 | -51.5034
0.040743 | -52.43
0.044447 | -52.6405
0.04815 | -54.2311
0.051854 | -55.2075
0.055558 | -56.8232
0.059262 | -57.2084
0.062966 | -57.0329
0.06667 | -58.2481
0.070374 | -59.206
0.074078 | -58.9802
0.077781 | -58.7976
0.081485 | -59.3046
0.085189 | -60.3557
0.088893 | -61.443
0.092597 | -62.0797
0.096301 | -62.3477
0.1 | -62.6671
0.10371 | -62.7164
0.10741 | -63.9273
0.11112 | -64.0347
0.11482 | -64.3019
0.11852 | -65.4003
0.12223 | -66.366
0.12593 | -66.0699
0.12964 | -66.1247
0.13334 | -65.3595
0.13704 | -65.2286
0.14075 | -65.5561
0.14445 | -65.4628
0.14816 | -67.041
0.15186 | -67.5411
0.15556 | -67.8283
0.15927 | -68.2796
0.16297 | -68.3017
0.16667 | -68.572
0.17038 | -69.0379
0.17408 | -68.9191
0.17779 | -69.0077
0.18149 | -69.6345
0.18519 | -69.1944
0.1889 | -69.1923
0.1926 | -69.8297
0.19631 | -70.2325
0.20001 | -70.2869
0.20371 | -71.0267
0.20742 | -71.3703
0.21112 | -72.025
0.21482 | -72.0403
0.21853 | -72.9766
0.22223 | -71.6433
0.22594 | -72.0714
0.22964 | -72.7416
0.23334 | -72.0111
0.23705 | -73.2243
0.24075 | -74.0026
0.24446 | -73.6862
0.24816 | -73.5535
0.25186 | -72.8978
0.25557 | -73.4953
0.25927 | -74.1263
0.26298 | -73.8407
0.26668 | -73.694
0.27038 | -73.5256
0.27409 | -72.9929
0.27779 | -73.683
0.28149 | -74.292
0.2852 | -75.2096
0.2889 | -74.7752
0.29261 | -74.0209
0.29631 | -75.4622
0.30001 | -76.0363
0.30372 | -75.7762
0.30742 | -74.9658
0.31113 | -74.3216
0.31483 | -75.7378
0.31853 | -76.2847
0.32224 | -75.9378
0.32594 | -76.6819
0.32965 | -76.5699
0.33335 | -75.8772
0.33705 | -76.1173
0.34076 | -76.533
0.34446 | -77.5356
0.34816 | -76.7974
0.35187 | -76.5179
0.35557 | -77.3581
0.35928 | -77.9518
0.36298 | -76.9102
0.36668 | -76.1341
0.37039 | -76.9904
0.37409 | -78.3098
0.37779 | -77.6888
0.3815 | -78.4944
0.3852 | -78.6726
0.38891 | -78.9034
0.39261 | -78.3363
0.39631 | -77.5543
0.40002 | -77.8848
0.40372 | -78.1454
0.40743 | -78.9847
0.41113 | -79.043
0.41483 | -77.6268
0.41854 | -77.5903
0.42224 | -78.1858
0.42595 | -78.8995
0.42965 | -79.2834
0.43335 | -79.3891
0.43706 | -79.4109
0.44076 | -78.4465
0.44446 | -79.054
0.44817 | -79.6509
0.45187 | -78.4355
0.45558 | -77.9222
0.45928 | -78.1799
0.46298 | -78.9588
0.46669 | -79.3096
0.47039 | -79.6741
0.4741 | -79.5946
0.4778 | -80.3513
0.4815 | -79.4937
0.48521 | -78.819
0.48891 | -80.2167
0.49262 | -81.0665
0.49632 | -81.7127
0.50002 | -80.7134
0.50373 | -80.4694
0.50743 | -80.2918
0.51114 | -80.2576
0.51484 | -81.1198
0.51854 | -81.8371
0.52225 | -82.1347
0.52595 | -82.13
0.52965 | -82.6029
0.53336 | -81.7793
0.53706 | -81.6169
0.54077 | -81.814
0.54447 | -81.8228
0.54817 | -81.4915
0.55188 | -80.9389
0.55558 | -80.516
0.55929 | -80.6951
0.56299 | -81.3335
0.56669 | -82.0469
0.5704 | -82.6119
0.5741 | -82.6362
0.57781 | -82.15
0.58151 | -81.8357
0.58521 | -81.8057
0.58892 | -82.2161
0.59262 | -82.7721
0.59632 | -82.9116
0.60003 | -82.6951
0.60373 | -82.2249
0.60744 | -81.8736
0.61114 | -81.7286
0.61484 | -82.0123
0.61855 | -82.6942
0.62225 | -82.8796
0.62596 | -82.3748
0.62966 | -82.1727
0.63336 | -82.6853
0.63707 | -83.2335
0.64077 | -83.3238
0.64447 | -83.5464
0.64818 | -83.6955
0.65188 | -83.4136
0.65559 | -83.338
0.65929 | -83.3396
0.66299 | -83.3244
0.6667 | -83.7053
0.6704 | -83.7699
0.6741 | -83.4506
0.67781 | -83.6348
0.68151 | -83.9572
0.68522 | -83.6153
0.68892 | -83.2024
0.69263 | -82.9645
0.69633 | -82.7344
0.70003 | -82.839
0.70374 | -83.4834
0.70744 | -84.0638
0.71114 | -84.3196
0.71485 | -84.369
0.71855 | -84.1545
0.72226 | -83.8859
0.72596 | -83.522
0.72966 | -83.6814
0.73337 | -84.2805
0.73707 | -84.7742
0.74077 | -85.0719
0.74448 | -85.2362
0.74818 | -85.0437
0.75189 | -84.915
0.75559 | -84.5089
0.75929 | -84.1627
0.763 | -84.5103
0.7667 | -84.9674
0.77041 | -85.0843
0.77411 | -85.227
0.77781 | -85.0209
0.78152 | -84.3934
0.78522 | -84.2548
0.78892 | -84.4287
0.79263 | -84.5646
0.79633 | -84.6728
0.80004 | -84.9661
0.80374 | -85.2983
0.80744 | -85.0004
0.81115 | -84.6508
0.81485 | -85.2068
0.81856 | -85.9216
0.82226 | -85.9052
0.82596 | -85.4684
0.82967 | -84.9418
0.83337 | -84.9202
0.83708 | -85.2429
0.84078 | -85.5368
0.84448 | -85.6339
0.84819 | -85.5877
0.85189 | -85.5697
0.85559 | -85.5848
0.8593 | -85.559
0.863 | -85.502
0.86671 | -85.6793
0.87041 | -85.9901
0.87411 | -86.0683
0.87782 | -86.2257
0.88152 | -86.6956
0.88523 | -87.1931
0.88893 | -87.3939
0.89263 | -86.8666
0.89634 | -86.7104
0.90004 | -86.8446
0.90375 | -86.588
0.90745 | -86.0268
0.91115 | -85.627
0.91486 | -85.7228
0.91856 | -86.0176
0.92227 | -86.418
0.92597 | -86.8634
0.92967 | -86.8153
0.93338 | -86.302
0.93708 | -86.2458
0.94078 | -86.7143
0.94449 | -87.1757
0.94819 | -87.0935
0.9519 | -86.8453
0.9556 | -86.7048
0.9593 | -86.8738
0.96301 | -87.0027
0.96671 | -86.9665
0.97042 | -86.8863
0.97412 | -87.0741
0.97782 | -87.1453
0.98153 | -87.0161
0.98523 | -86.6498
0.98894 | -86.2406
0.99264 | -86.213
0.99634 | -86.1746
1 | -85.9058
1.0037 | -86.0019
1.0075 | -86.7004
1.0112 | -87.3508
1.0149 | -87.5973
1.0186 | -87.4316
1.0223 | -87.3781
1.026 | -87.5948
1.0297 | -87.8019
1.0334 | -88.3262
1.0371 | -88.7253
1.0408 | -88.4854
1.0445 | -87.725
1.0482 | -87.0404
1.0519 | -87.1125
1.0556 | -87.5638
1.0593 | -87.8533
1.063 | -88.0806
1.0667 | -87.7491
1.0704 | -87.5847
1.0741 | -87.5585
1.0778 | -87.5597
1.0815 | -87.4949
1.0852 | -87.361
1.0889 | -87.2885
1.0926 | -87.5083
1.0963 | -87.9423
1.1001 | -88.0516
1.1038 | -87.6988
1.1075 | -87.4406
1.1112 | -87.5659
1.1149 | -87.8495
1.1186 | -88.0239
1.1223 | -88.1436
1.126 | -88.3069
1.1297 | -88.4591
1.1334 | -88.5942
1.1371 | -88.7429
1.1408 | -88.8854
1.1445 | -88.918
1.1482 | -88.6242
1.1519 | -88.048
1.1556 | -87.5941
1.1593 | -87.5694
1.163 | -87.9372
1.1667 | -88.3893
1.1704 | -88.6341
1.1741 | -88.6424
1.1778 | -88.6612
1.1815 | -88.879
1.1852 | -89.1748
1.1889 | -89.3886
1.1926 | -89.4489
1.1964 | -89.3903
1.2001 | -89.3571
1.2038 | -89.3063
1.2075 | -89.1683
1.2112 | -89.0241
1.2149 | -88.9166
1.2186 | -88.7981
1.2223 | -88.6416
1.226 | -88.5421
1.2297 | -88.6388
1.2334 | -88.9752
1.2371 | -89.4884
1.2408 | -89.9494
1.2445 | -90.0067
1.2482 | -89.65
1.2519 | -89.2822
1.2556 | -89.1978
1.2593 | -89.3274
1.263 | -89.369
1.2667 | -89.1604
1.2704 | -88.7941
1.2741 | -88.4531
1.2778 | -88.3182
1.2815 | -88.4183
1.2852 | -88.7249
1.2889 | -89.1628
1.2927 | -89.446
1.2964 | -89.5216
1.3001 | -89.6923
1.3038 | -89.9272
1.3075 | -90.0106
1.3112 | -89.9802
1.3149 | -90.0888
1.3186 | -90.4337
1.3223 | -90.7408
1.326 | -90.6648
1.3297 | -90.373
1.3334 | -90.2159
1.3371 | -90.1708
1.3408 | -89.9837
1.3445 | -89.6304
1.3482 | -89.3404
1.3519 | -89.2856
1.3556 | -89.4135
1.3593 | -89.6819
1.363 | -90.1162
1.3667 | -90.5636
1.3704 | -90.6896
1.3741 | -90.3136
1.3778 | -89.7103
1.3815 | -89.3634
1.3852 | -89.3762
1.389 | -89.465
1.3927 | -89.495
1.3964 | -89.6436
1.4001 | -89.9169
1.4038 | -90.0923
1.4075 | -90.0766
1.4112 | -89.9217
1.4149 | -89.7595
1.4186 | -89.8024
1.4223 | -90.0871
1.426 | -90.382
1.4297 | -90.4845
1.4334 | -90.4927
1.4371 | -90.526
1.4408 | -90.5841
1.4445 | -90.7754
1.4482 | -91.1453
1.4519 | -91.3277
1.4556 | -90.9612
1.4593 | -90.3958
1.463 | -90.0809
1.4667 | -89.9572
1.4704 | -89.8757
1.4741 | -90.0109
1.4778 | -90.4291
1.4815 | -90.7897
1.4853 | -90.9205
1.489 | -91.1024
1.4927 | -91.3676
1.4964 | -91.4333
1.5001 | -91.2832
1.5038 | -91.0163
1.5075 | -90.6903
1.5112 | -90.5271
1.5149 | -90.6277
1.5186 | -90.7842
1.5223 | -90.797
1.526 | -90.7386
1.5297 | -90.734
1.5334 | -90.7252
1.5371 | -90.6829
1.5408 | -90.7521
1.5445 | -90.9389
1.5482 | -91.0371
1.5519 | -90.9305
1.5556 | -90.8042
1.5593 | -90.8883
1.563 | -91.145
1.5667 | -91.4091
1.5704 | -91.6281
1.5741 | -91.7558
1.5779 | -91.7589
1.5816 | -91.7638
1.5853 | -91.7662
1.589 | -91.576
1.5927 | -91.4178
1.5964 | -91.5444
1.6001 | -91.5659
1.6038 | -91.4961
1.6075 | -91.4035
1.6112 | -91.3976
1.6149 | -91.5368
1.6186 | -91.7841
1.6223 | -92
1.626 | -92.049
1.6297 | -91.8911
1.6334 | -91.6283
1.6371 | -91.4295
1.6408 | -91.4388
1.6445 | -91.6887
1.6482 | -92.0618
1.6519 | -92.3644
1.6556 | -92.4573
1.6593 | -92.369
1.663 | -92.2277
1.6667 | -92.1473
1.6704 | -92.1678
1.6742 | -92.2429
1.6779 | -92.2437
1.6816 | -92.092
1.6853 | -91.8715
1.689 | -91.7165
1.6927 | -91.7125
1.6964 | -91.8317
1.7001 | -91.949
1.7038 | -91.949
1.7075 | -91.8291
1.7112 | -91.7034
1.7149 | -91.6175
1.7186 | -91.5552
1.7223 | -91.4797
1.726 | -91.4061
1.7297 | -91.4256
1.7334 | -91.6029
1.7371 | -91.8883
1.7408 | -92.1164
1.7445 | -92.1825
1.7482 | -92.1787
1.7519 | -92.2384
1.7556 | -92.4358
1.7593 | -92.6723
1.763 | -92.7822
1.7667 | -92.7237
1.7705 | -92.5885
1.7742 | -92.4921
1.7779 | -92.5005
1.7816 | -92.6166
1.7853 | -92.7524
1.789 | -92.7867
1.7927 | -92.7142
1.7964 | -92.5986
1.8001 | -92.4706
1.8038 | -92.3146
1.8075 | -92.1396
1.8112 | -92.0082
1.8149 | -92.0213
1.8186 | -92.2253
1.8223 | -92.5608
1.826 | -92.8741
1.8297 | -93.0108
1.8334 | -92.93
1.8371 | -92.7352
1.8408 | -92.572
1.8445 | -92.5109
1.8482 | -92.5524
1.8519 | -92.7046
1.8556 | -92.9222
1.8593 | -93.0836
1.863 | -93.0897
1.8668 | -93.0157
1.8705 | -92.9872
1.8742 | -93.0546
1.8779 | -93.1828
1.8816 | -93.2779
1.8853 | -93.2898
1.889 | -93.245
1.8927 | -93.1922
1.8964 | -93.1531
1.9001 | -93.135
1.9038 | -93.1632
1.9075 | -93.2212
1.9112 | -93.2712
1.9149 | -93.2571
1.9186 | -93.1845
1.9223 | -93.1375
1.926 | -93.2132
1.9297 | -93.3766
1.9334 | -93.4363
1.9371 | -93.2634
1.9408 | -92.9789
1.9445 | -92.7869
1.9482 | -92.7871
1.9519 | -92.9304
1.9556 | -93.1251
1.9593 | -93.2831
1.9631 | -93.3666
1.9668 | -93.4016
1.9705 | -93.4387
1.9742 | -93.4868
1.9779 | -93.5514
1.9816 | -93.6454
1.9853 | -93.7867
1.989 | -93.9664
1.9927 | -94.1291
1.9964 | -94.2135
2.0001 | -94.1733
2.0038 | -94.0083
2.0075 | -93.7555
2.0112 | -93.5435
2.0149 | -93.4948
2.0186 | -93.5583
2.0223 | -93.5435
2.026 | -93.3759
2.0297 | -93.1448
2.0334 | -92.9802
2.0371 | -92.9374
2.0408 | -93.0096
2.0445 | -93.1342
2.0482 | -93.265
2.0519 | -93.3712
2.0557 | -93.4088
2.0594 | -93.3701
2.0631 | -93.3215
2.0668 | -93.3166
2.0705 | -93.3963
2.0742 | -93.5858
2.0779 | -93.7961
2.0816 | -93.9112
2.0853 | -93.9087
2.089 | -93.878
2.0927 | -93.8772
2.0964 | -93.8966
2.1001 | -93.9127
2.1038 | -93.882
2.1075 | -93.8011
2.1112 | -93.7862
2.1149 | -93.8802
2.1186 | -94.0145
2.1223 | -94.085
2.126 | -94.0867
2.1297 | -94.0809
2.1334 | -94.0779
2.1371 | -94.0496
2.1408 | -94.0031
2.1445 | -93.9385
2.1482 | -93.8813
2.152 | -93.8704
2.1557 | -93.9273
2.1594 | -94.0326
2.1631 | -94.1496
2.1668 | -94.2367
2.1705 | -94.2666
2.1742 | -94.2196
2.1779 | -94.1036
2.1816 | -93.9563
2.1853 | -93.8389
2.189 | -93.8131
2.1927 | -93.9349
2.1964 | -94.1927
2.2001 | -94.5101
2.2038 | -94.7787
2.2075 | -94.9074
2.2112 | -94.8757
2.2149 | -94.7423
2.2186 | -94.5738
2.2223 | -94.4255
2.226 | -94.3108
2.2297 | -94.2364
2.2334 | -94.211
2.2371 | -94.2268
2.2408 | -94.2452
2.2445 | -94.2203
2.2483 | -94.1678
2.252 | -94.1242
2.2557 | -94.1221
2.2594 | -94.1798
2.2631 | -94.2434
2.2668 | -94.2524
2.2705 | -94.1804
2.2742 | -94.045
2.2779 | -93.9074
2.2816 | -93.8258
2.2853 | -93.8463
2.289 | -93.9651
2.2927 | -94.1313
2.2964 | -94.2779
2.3001 | -94.3534
2.3038 | -94.381
2.3075 | -94.3872
2.3112 | -94.3879
2.3149 | -94.3861
2.3186 | -94.3734
2.3223 | -94.3642
2.326 | -94.3645
2.3297 | -94.3699
2.3334 | -94.3923
2.3371 | -94.4299
2.3408 | -94.5049
2.3446 | -94.6268
2.3483 | -94.7775
2.352 | -94.9184
2.3557 | -95.0387
2.3594 | -95.1452
2.3631 | -95.2407
2.3668 | -95.3464
2.3705 | -95.4209
2.3742 | -95.4012
2.3779 | -95.3013
2.3816 | -95.1558
2.3853 | -95.0315
2.389 | -94.9592
2.3927 | -94.9216
2.3964 | -94.8818
2.4001 | -94.8227
2.4038 | -94.7408
2.4075 | -94.6422
2.4112 | -94.5367
2.4149 | -94.4258
2.4186 | -94.3412
2.4223 | -94.3178
2.426 | -94.3824
2.4297 | -94.5296
2.4334 | -94.7264
2.4371 | -94.9144
2.4409 | -95.0349
2.4446 | -95.093
2.4483 | -95.0949
2.452 | -95.082
2.4557 | -95.0773
2.4594 | -95.0762
2.4631 | -95.0852
2.4668 | -95.0967
2.4705 | -95.1215
2.4742 | -95.1519
2.4779 | -95.1444
2.4816 | -95.0605
2.4853 | -94.8989
2.489 | -94.7364
2.4927 | -94.653
2.4964 | -94.6919
2.5001 | -94.8403
2.5038 | -95.0666
2.5075 | -95.3236
2.5112 | -95.5576
2.5149 | -95.7057
2.5186 | -95.7449
2.5223 | -95.6646
2.526 | -95.5019
2.5297 | -95.3159
2.5335 | -95.1778
2.5372 | -95.1592
2.5409 | -95.2513
2.5446 | -95.3468
2.5483 | -95.3294
2.552 | -95.1909
2.5557 | -95.0169
2.5594 | -94.9154
2.5631 | -94.942
2.5668 | -95.079
2.5705 | -95.2668
2.5742 | -95.4344
2.5779 | -95.5285
2.5816 | -95.5493
2.5853 | -95.5219
2.589 | -95.4851
2.5927 | -95.4767
2.5964 | -95.5194
2.6001 | -95.616
2.6038 | -95.7385
2.6075 | -95.8313
2.6112 | -95.8361
2.6149 | -95.7277
2.6186 | -95.5423
2.6223 | -95.3371
2.626 | -95.1687
2.6298 | -95.0684
2.6335 | -95.0343
2.6372 | -95.024
2.6409 | -94.9984
2.6446 | -94.9727
2.6483 | -94.9803
2.652 | -95.0483
2.6557 | -95.1559
2.6594 | -95.2456
2.6631 | -95.3041
2.6668 | -95.3812
2.6705 | -95.5858
2.6742 | -95.779
2.6779 | -95.9652
2.6816 | -96.1199
2.6853 | -96.2092
2.689 | -96.2277
2.6927 | -96.1972
2.6964 | -96.1495
2.7001 | -96.1223
2.7038 | -96.1606
2.7075 | -96.2711
2.7112 | -96.4322
2.7149 | -96.6101
2.7186 | -96.771
2.7223 | -96.9021
2.7261 | -96.9854
2.7298 | -96.9902
2.7335 | -96.918
2.7372 | -96.7763
2.7409 | -96.5795
2.7446 | -96.3675
2.7483 | -96.1821
2.752 | -96.0553
2.7557 | -96.0166
2.7594 | -96.0761
2.7631 | -96.2267
2.7668 | -96.4241
2.7705 | -96.5963
2.7742 | -96.6824
2.7779 | -96.6649
2.7816 | -96.5691
2.7853 | -96.4377
2.789 | -96.2924
2.7927 | -96.15
2.7964 | -96.0258
2.8001 | -95.9292
2.8038 | -95.8568
2.8075 | -95.8052
2.8112 | -95.7737
2.8149 | -95.7756
2.8186 | -95.8327
2.8224 | -95.9307
2.8261 | -96.0463
2.8298 | -96.1577
2.8335 | -96.2606
2.8372 | -96.3791
2.8409 | -96.5015
2.8446 | -96.6179
2.8483 | -96.6879
2.852 | -96.6783
2.8557 | -96.5855
2.8594 | -96.4263
2.8631 | -96.251
2.8668 | -96.1132
2.8705 | -96.0587
2.8742 | -96.0955
2.8779 | -96.2078
2.8816 | -96.3556
2.8853 | -96.4711
2.889 | -96.518
2.8927 | -96.4933
2.8964 | -96.422
2.9001 | -96.3372
2.9038 | -96.2636
2.9075 | -96.2132
2.9112 | -96.196
2.915 | -96.2215
2.9187 | -96.2936
2.9224 | -96.3774
2.9261 | -96.4186
2.9298 | -96.3835
2.9335 | -96.2635
2.9372 | -96.0912
2.9409 | -95.9247
2.9446 | -95.8277
2.9483 | -95.8212
2.952 | -95.8998
2.9557 | -96.0465
2.9594 | -96.2189
2.9631 | -96.3799
2.9668 | -96.5068
2.9705 | -96.6047
2.9742 | -96.7047
2.9779 | -96.823
2.9816 | -96.9606
2.9853 | -97.0807
2.989 | -97.1354
2.9927 | -97.0968
2.9964 | -96.9926
3.0001 | -96.8415
3.0038 | -96.6747
3.0075 | -96.5472
3.0113 | -96.4804
3.015 | -96.4778
3.0187 | -96.5299
3.0224 | -96.6179
3.0261 | -96.7453
3.0298 | -96.9076
3.0335 | -97.0689
3.0372 | -97.2021
3.0409 | -97.2849
3.0446 | -97.293
3.0483 | -97.2549
3.052 | -97.1824
3.0557 | -97.0693
3.0594 | -96.9526
3.0631 | -96.8637
3.0668 | -96.818
3.0705 | -96.8207
3.0742 | -96.8398
3.0779 | -96.8504
3.0816 | -96.832
3.0853 | -96.7743
3.089 | -96.7083
3.0927 | -96.6692
3.0964 | -96.6823
3.1001 | -96.7364
3.1038 | -96.7962
3.1076 | -96.8112
3.1113 | -96.7551
3.115 | -96.6695
3.1187 | -96.6145
3.1224 | -96.6467
3.1261 | -96.7824
3.1298 | -97.0096
3.1335 | -97.2705
3.1372 | -97.4846
3.1409 | -97.6056
3.1446 | -97.5971
3.1483 | -97.5004
3.152 | -97.3599
3.1557 | -97.2178
3.1594 | -97.1163
3.1631 | -97.0635
3.1668 | -97.0595
3.1705 | -97.0762
3.1742 | -97.0786
3.1779 | -97.0535
3.1816 | -97.0179
3.1853 | -96.9849
3.189 | -96.9518
3.1927 | -96.943
3.1964 | -96.948
3.2001 | -96.9622
3.2039 | -96.9791
3.2076 | -97.0047
3.2113 | -97.0452
3.215 | -97.0954
3.2187 | -97.1612
3.2224 | -97.2416
3.2261 | -97.3222
3.2298 | -97.3965
3.2335 | -97.4521
3.2372 | -97.4952
3.2409 | -97.5406
3.2446 | -97.5909
3.2483 | -97.6651
3.252 | -97.7574
3.2557 | -97.8467
3.2594 | -97.9152
3.2631 | -97.9343
3.2668 | -97.903
3.2705 | -97.8353
3.2742 | -97.744
3.2779 | -97.6512
3.2816 | -97.5748
3.2853 | -97.516
3.289 | -97.4963
3.2927 | -97.5179
3.2964 | -97.5726
3.3002 | -97.6497
3.3039 | -97.7315
3.3076 | -97.8068
3.3113 | -97.8802
3.315 | -97.9341
3.3187 | -97.9732
3.3224 | -97.99
3.3261 | -98.002
3.3298 | -98.0056
3.3335 | -98.0149
3.3372 | -98.026
3.3409 | -98.0335
3.3446 | -98.0211
3.3483 | -97.9939
3.352 | -97.9512
3.3557 | -97.8967
3.3594 | -97.8367
3.3631 | -97.7614
3.3668 | -97.6675
3.3705 | -97.5875
3.3742 | -97.5259
3.3779 | -97.4954
3.3816 | -97.4996
3.3853 | -97.5559
3.389 | -97.6456
3.3927 | -97.7562
3.3965 | -97.8653
3.4002 | -97.9519
3.4039 | -98.0028
3.4076 | -98.0176
3.4113 | -97.9953
3.415 | -97.9539
3.4187 | -97.9089
3.4224 | -97.8663
3.4261 | -97.834
3.4298 | -97.82
3.4335 | -97.8138
3.4372 | -97.8082
3.4409 | -97.7892
3.4446 | -97.7446
3.4483 | -97.6681
3.452 | -97.5703
3.4557 | -97.4611
3.4594 | -97.3538
3.4631 | -97.2708
3.4668 | -97.2286
3.4705 | -97.2246
3.4742 | -97.272
3.4779 | -97.3687
3.4816 | -97.4912
3.4853 | -97.6249
3.4891 | -97.7659
3.4928 | -97.8895
3.4965 | -97.9985
3.5002 | -98.0927
3.5039 | -98.1676
3.5076 | -98.2213
3.5113 | -98.264
3.515 | -98.3062
3.5187 | -98.358
3.5224 | -98.4267
3.5261 | -98.5076
3.5298 | -98.5782
3.5335 | -98.6291
3.5372 | -98.669
3.5409 | -98.6912
3.5446 | -98.6963
3.5483 | -98.7184
3.552 | -98.7754
3.5557 | -98.8578
3.5594 | -98.9399
3.5631 | -99.001
3.5668 | -99.0223
3.5705 | -98.9903
3.5742 | -98.9086
3.5779 | -98.781
3.5816 | -98.6137
3.5854 | -98.4259
3.5891 | -98.251
3.5928 | -98.1137
3.5965 | -98.0431
3.6002 | -98.0464
3.6039 | -98.108
3.6076 | -98.1935
3.6113 | -98.2859
3.615 | -98.3588
3.6187 | -98.404
3.6224 | -98.4147
3.6261 | -98.3825
3.6298 | -98.3147
3.6335 | -98.2201
3.6372 | -98.1196
3.6409 | -98.0309
3.6446 | -97.9678
3.6483 | -97.9375
3.652 | -97.9339
3.6557 | -97.954
3.6594 | -97.9991
3.6631 | -98.0728
3.6668 | -98.1687
3.6705 | -98.2769
3.6742 | -98.3583
3.6779 | -98.3915
3.6817 | -98.3646
3.6854 | -98.2763
3.6891 | -98.1641
3.6928 | -98.0668
3.6965 | -98.0131
3.7002 | -98.0175
3.7039 | -98.0922
3.7076 | -98.2162
3.7113 | -98.3646
3.715 | -98.5076
3.7187 | -98.6085
3.7224 | -98.6562
3.7261 | -98.6531
3.7298 | -98.6187
3.7335 | -98.5794
3.7372 | -98.5745
3.7409 | -98.5672
3.7446 | -98.5762
3.7483 | -98.5993
3.752 | -98.6227
3.7557 | -98.6448
3.7594 | -98.6617
3.7631 | -98.6517
3.7668 | -98.6125
3.7705 | -98.5513
3.7742 | -98.486
3.778 | -98.4425
3.7817 | -98.4195
3.7854 | -98.4218
3.7891 | -98.4619
3.7928 | -98.5232
3.7965 | -98.6008
3.8002 | -98.6705
3.8039 | -98.7222
3.8076 | -98.7344
3.8113 | -98.7108
3.815 | -98.6673
3.8187 | -98.605
3.8224 | -98.5444
3.8261 | -98.4951
3.8298 | -98.4681
3.8335 | -98.4666
3.8372 | -98.4899
3.8409 | -98.5343
3.8446 | -98.5967
3.8483 | -98.6705
3.852 | -98.7411
3.8557 | -98.8116
3.8594 | -98.8801
3.8631 | -98.9502
3.8668 | -99.0092
3.8706 | -99.0572
3.8743 | -99.0979
3.878 | -99.1387
3.8817 | -99.1808
3.8854 | -99.2186
3.8891 | -99.2507
3.8928 | -99.2618
3.8965 | -99.2522
3.9002 | -99.2348
3.9039 | -99.2129
3.9076 | -99.1978
3.9113 | -99.1974
3.915 | -99.2149
3.9187 | -99.2342
3.9224 | -99.2471
3.9261 | -99.2531
3.9298 | -99.2553
3.9335 | -99.2528
3.9372 | -99.2365
3.9409 | -99.2159
3.9446 | -99.2033
3.9483 | -99.1921
3.952 | -99.2006
3.9557 | -99.2224
3.9594 | -99.248
3.9631 | -99.2651
3.9668 | -99.2696
3.9706 | -99.2343
3.9743 | -99.1476
3.978 | -99.0134
3.9817 | -98.8588
3.9854 | -98.7124
3.9891 | -98.5818
3.9928 | -98.4852
3.9965 | -98.4403
4.0002 | -98.4478
4.0039 | -98.5013
4.0076 | -98.5862
4.0113 | -98.6926
4.015 | -98.8029
4.0187 | -98.9028
4.0224 | -98.9809
4.0261 | -99.0311
4.0298 | -99.05
4.0335 | -99.0622
4.0372 | -99.0781
4.0409 | -99.111
4.0446 | -99.1685
4.0483 | -99.2483
4.052 | -99.3507
4.0557 | -99.4592
4.0594 | -99.5535
4.0632 | -99.6274
4.0669 | -99.6751
4.0706 | -99.7
4.0743 | -99.6899
4.078 | -99.6459
4.0817 | -99.576
4.0854 | -99.4868
4.0891 | -99.3822
4.0928 | -99.2722
4.0965 | -99.152
4.1002 | -99.0432
4.1039 | -98.9544
4.1076 | -98.8935
4.1113 | -98.8666
4.115 | -98.8923
4.1187 | -98.9676
4.1224 | -99.087
4.1261 | -99.2369
4.1298 | -99.3845
4.1335 | -99.5084
4.1372 | -99.5925
4.1409 | -99.6335
4.1446 | -99.6446
4.1483 | -99.6307
4.152 | -99.6067
4.1557 | -99.5749
4.1595 | -99.5459
4.1632 | -99.5181
4.1669 | -99.4848
4.1706 | -99.4417
4.1743 | -99.3895
4.178 | -99.3308
4.1817 | -99.2715
4.1854 | -99.2152
4.1891 | -99.1708
4.1928 | -99.147
4.1965 | -99.1385
4.2002 | -99.1404
4.2039 | -99.156
4.2076 | -99.1743
4.2113 | -99.1928
4.215 | -99.2072
4.2187 | -99.2206
4.2224 | -99.2448
4.2261 | -99.2774
4.2298 | -99.3383
4.2335 | -99.4164
4.2372 | -99.4939
4.2409 | -99.5633
4.2446 | -99.6213
4.2483 | -99.6634
4.252 | -99.7014
4.2558 | -99.7379
4.2595 | -99.7765
4.2632 | -99.8187
4.2669 | -99.86
4.2706 | -99.919
4.2743 | -99.9801
4.278 | -100.025
4.2817 | -100.0661
4.2854 | -100.0862
4.2891 | -100.0739
4.2928 | -100.0423
4.2965 | -99.9867
4.3002 | -99.916
4.3039 | -99.8403
4.3076 | -99.7649
4.3113 | -99.6997
4.315 | -99.6514
4.3187 | -99.6216
4.3224 | -99.6169
4.3261 | -99.6405
4.3298 | -99.6912
4.3335 | -99.7618
4.3372 | -99.842
4.3409 | -99.932
4.3446 | -100.0212
4.3483 | -100.1001
4.3521 | -100.1589
4.3558 | -100.2048
4.3595 | -100.2295
4.3632 | -100.2324
4.3669 | -100.2154
4.3706 | -100.189
4.3743 | -100.1503
4.378 | -100.0999
4.3817 | -100.048
4.3854 | -99.9948
4.3891 | -99.9393
4.3928 | -99.8828
4.3965 | -99.817
4.4002 | -99.7524
4.4039 | -99.6884
4.4076 | -99.6295
4.4113 | -99.5718
4.415 | -99.5262
4.4187 | -99.4894
4.4224 | -99.4643
4.4261 | -99.455
4.4298 | -99.4536
4.4335 | -99.4633
4.4372 | -99.4793
4.4409 | -99.5034
4.4447 | -99.5426
4.4484 | -99.5935
4.4521 | -99.6519
4.4558 | -99.713
4.4595 | -99.7784
4.4632 | -99.8452
4.4669 | -99.9096
4.4706 | -99.97
4.4743 | -100.0386
4.478 | -100.1164
4.4817 | -100.1918
4.4854 | -100.2618
4.4891 | -100.3185
4.4928 | -100.352
4.4965 | -100.3622
4.5002 | -100.3635
4.5039 | -100.3485
4.5076 | -100.337
4.5113 | -100.321
4.515 | -100.2954
4.5187 | -100.2601
4.5224 | -100.2172
4.5261 | -100.1675
4.5298 | -100.1029
4.5335 | -100.0363
4.5372 | -99.976
4.541 | -99.9255
4.5447 | -99.8918
4.5484 | -99.8788
4.5521 | -99.8896
4.5558 | -99.9067
4.5595 | -99.9291
4.5632 | -99.9555
4.5669 | -99.977
4.5706 | -99.9818
4.5743 | -99.9802
4.578 | -99.9721
4.5817 | -99.9621
4.5854 | -99.9495
4.5891 | -99.9377
4.5928 | -99.9222
4.5965 | -99.9081
4.6002 | -99.8995
4.6039 | -99.8953
4.6076 | -99.8961
4.6113 | -99.9025
4.615 | -99.91
4.6187 | -99.922
4.6224 | -99.9324
4.6261 | -99.9335
4.6298 | -99.9266
4.6335 | -99.9081
4.6373 | -99.8786
4.641 | -99.8502
4.6447 | -99.8198
4.6484 | -99.8029
4.6521 | -99.8058
4.6558 | -99.8297
4.6595 | -99.8735
4.6632 | -99.9335
4.6669 | -100.0081
4.6706 | -100.091
4.6743 | -100.1691
4.678 | -100.2474
4.6817 | -100.3208
4.6854 | -100.3887
4.6891 | -100.438
4.6928 | -100.4767
4.6965 | -100.5082
4.7002 | -100.5295
4.7039 | -100.5501
4.7076 | -100.5762
4.7113 | -100.617
4.715 | -100.6599
4.7187 | -100.6935
4.7224 | -100.7121
4.7261 | -100.711
4.7298 | -100.681
4.7336 | -100.63
4.7373 | -100.5595
4.741 | -100.4856
4.7447 | -100.4243
4.7484 | -100.3684
4.7521 | -100.326
4.7558 | -100.2939
4.7595 | -100.2807
4.7632 | -100.2759
4.7669 | -100.2752
4.7706 | -100.2793
4.7743 | -100.2809
4.778 | -100.2774
4.7817 | -100.2675
4.7854 | -100.2723
4.7891 | -100.3004
4.7928 | -100.3545
4.7965 | -100.4352
4.8002 | -100.5422
4.8039 | -100.6576
4.8076 | -100.7425
4.8113 | -100.8153
4.815 | -100.8698
4.8187 | -100.9041
4.8224 | -100.921
4.8261 | -100.9204
4.8299 | -100.9078
4.8336 | -100.8839
4.8373 | -100.8497
4.841 | -100.8169
4.8447 | -100.7923
4.8484 | -100.782
4.8521 | -100.7831
4.8558 | -100.7927
4.8595 | -100.805
4.8632 | -100.8178
4.8669 | -100.8349
4.8706 | -100.8431
4.8743 | -100.8514
4.878 | -100.8642
4.8817 | -100.8775
4.8854 | -100.8931
4.8891 | -100.9141
4.8928 | -100.9334
4.8965 | -100.9581
4.9002 | -100.9888
4.9039 | -101.0243
4.9076 | -101.0665
4.9113 | -101.1035
4.915 | -101.1396
4.9187 | -101.1663
4.9225 | -101.1725
4.9262 | -101.169
4.9299 | -101.1539
4.9336 | -101.1153
4.9373 | -101.0621
4.941 | -100.9956
4.9447 | -100.9169
4.9484 | -100.8405
4.9521 | -100.7654
4.9558 | -100.7021
4.9595 | -100.6581
4.9632 | -100.6298
4.9669 | -100.6271
4.9706 | -100.6463
4.9743 | -100.6828
4.978 | -100.7364
4.9817 | -100.8029
4.9854 | -100.87
4.9891 | -100.9378
4.9928 | -101.0073
4.9965 | -101.0716
5.0002 | -101.124
5.0039 | -101.1509
5.0076 | -101.1586
5.0113 | -101.1532
5.015 | -101.1132
5.0188 | -101.0503
5.0225 | -100.9669
5.0262 | -100.8701
5.0299 | -100.7837
5.0336 | -100.7159
5.0373 | -100.6746
5.041 | -100.668
5.0447 | -100.6894
5.0484 | -100.7403
5.0521 | -100.8233
5.0558 | -100.9283
5.0595 | -101.042
5.0632 | -101.1588
5.0669 | -101.2656
5.0706 | -101.3435
5.0743 | -101.3979
5.078 | -101.4294
5.0817 | -101.435
5.0854 | -101.4105
5.0891 | -101.3622
5.0928 | -101.3052
5.0965 | -101.2387
5.1002 | -101.1825
5.1039 | -101.1375
5.1076 | -101.0998
5.1113 | -101.0769
5.1151 | -101.0656
5.1188 | -101.067
5.1225 | -101.0793
5.1262 | -101.0988
5.1299 | -101.1198
5.1336 | -101.1484
5.1373 | -101.1853
5.141 | -101.2296
5.1447 | -101.2722
5.1484 | -101.3185
5.1521 | -101.3646
5.1558 | -101.4012
5.1595 | -101.4271
5.1632 | -101.4392
5.1669 | -101.4364
5.1706 | -101.4161
5.1743 | -101.3766
5.178 | -101.3156
5.1817 | -101.238
5.1854 | -101.1532
5.1891 | -101.0712
5.1928 | -101.0039
5.1965 | -100.9417
5.2002 | -100.8914
5.2039 | -100.8646
5.2076 | -100.854
5.2114 | -100.849
5.2151 | -100.8519
5.2188 | -100.864
5.2225 | -100.8868
5.2262 | -100.9201
5.2299 | -100.9534
5.2336 | -100.9957
5.2373 | -101.0489
5.241 | -101.1164
5.2447 | -101.1933
5.2484 | -101.2687
5.2521 | -101.3287
5.2558 | -101.382
5.2595 | -101.4168
5.2632 | -101.4242
5.2669 | -101.4093
5.2706 | -101.3686
5.2743 | -101.3196
5.278 | -101.2596
5.2817 | -101.1944
5.2854 | -101.1366
5.2891 | -101.1014
5.2928 | -101.0856
5.2965 | -101.0905
5.3002 | -101.117
5.3039 | -101.156
5.3077 | -101.2079
5.3114 | -101.2638
5.3151 | -101.2989
5.3188 | -101.3224
5.3225 | -101.3384
5.3262 | -101.3542
5.3299 | -101.3761
5.3336 | -101.3937
5.3373 | -101.4234
5.341 | -101.4595
5.3447 | -101.4924
5.3484 | -101.5208
5.3521 | -101.5523
5.3558 | -101.5834
5.3595 | -101.6132
5.3632 | -101.6468
5.3669 | -101.6884
5.3706 | -101.731
5.3743 | -101.7714
5.378 | -101.8071
5.3817 | -101.8373
5.3854 | -101.8562
5.3891 | -101.8624
5.3928 | -101.8448
5.3965 | -101.8138
5.4002 | -101.7568
5.404 | -101.6786
5.4077 | -101.6024
5.4114 | -101.5256
5.4151 | -101.4465
5.4188 | -101.3694
5.4225 | -101.3071
5.4262 | -101.2613
5.4299 | -101.2223
5.4336 | -101.1959
5.4373 | -101.1729
5.441 | -101.1541
5.4447 | -101.1483
5.4484 | -101.1428
5.4521 | -101.1397
5.4558 | -101.1425
5.4595 | -101.1505
5.4632 | -101.1615
5.4669 | -101.172
5.4706 | -101.1851
5.4743 | -101.197
5.478 | -101.2034
5.4817 | -101.2098
5.4854 | -101.2177
5.4891 | -101.2285
5.4928 | -101.2333
5.4966 | -101.2405
5.5003 | -101.2546
5.504 | -101.2657
5.5077 | -101.27
5.5114 | -101.2727
5.5151 | -101.2769
5.5188 | -101.2881
5.5225 | -101.3083
5.5262 | -101.3331
5.5299 | -101.358
5.5336 | -101.3901
5.5373 | -101.4209
5.541 | -101.4504
5.5447 | -101.4839
5.5484 | -101.5208
5.5521 | -101.5694
5.5558 | -101.6254
5.5595 | -101.6869
5.5632 | -101.7543
5.5669 | -101.8144
5.5706 | -101.8704
5.5743 | -101.9138
5.578 | -101.9453
5.5817 | -101.9714
5.5854 | -101.9915
5.5891 | -102.0094
5.5929 | -102.016
5.5966 | -102.0057
5.6003 | -101.98
5.604 | -101.9362
5.6077 | -101.8844
5.6114 | -101.8215
5.6151 | -101.7585
5.6188 | -101.6937
5.6225 | -101.6334
5.6262 | -101.573
5.6299 | -101.5207
5.6336 | -101.4823
5.6373 | -101.4484
5.641 | -101.4233
5.6447 | -101.415
5.6484 | -101.4268
5.6521 | -101.4538
5.6558 | -101.4843
5.6595 | -101.5323
5.6632 | -101.5828
5.6669 | -101.6259
5.6706 | -101.6646
5.6743 | -101.6983
5.678 | -101.7285
5.6817 | -101.7506
5.6854 | -101.7692
5.6892 | -101.7721
5.6929 | -101.7648
5.6966 | -101.7632
5.7003 | -101.7557
5.704 | -101.7504
5.7077 | -101.7483
5.7114 | -101.7475
5.7151 | -101.745
5.7188 | -101.7435
5.7225 | -101.7479
5.7262 | -101.7644
5.7299 | -101.7832
5.7336 | -101.7929
5.7373 | -101.7987
5.741 | -101.8069
5.7447 | -101.7994
5.7484 | -101.7825
5.7521 | -101.7648
5.7558 | -101.742
5.7595 | -101.7128
5.7632 | -101.692
5.7669 | -101.6818
5.7706 | -101.6746
5.7743 | -101.6833
5.778 | -101.7019
5.7817 | -101.7239
5.7855 | -101.748
5.7892 | -101.7746
5.7929 | -101.8014
5.7966 | -101.8223
5.8003 | -101.8347
5.804 | -101.8417
5.8077 | -101.8406
5.8114 | -101.8204
5.8151 | -101.7892
5.8188 | -101.7494
5.8225 | -101.6998
5.8262 | -101.6393
5.8299 | -101.5778
5.8336 | -101.5234
5.8373 | -101.474
5.841 | -101.4327
5.8447 | -101.4056
5.8484 | -101.392
5.8521 | -101.4015
5.8558 | -101.4319
5.8595 | -101.4709
5.8632 | -101.5245
5.8669 | -101.5817
5.8706 | -101.6489
5.8743 | -101.6928
5.878 | -101.7352
5.8818 | -101.7696
5.8855 | -101.7973
5.8892 | -101.8156
5.8929 | -101.8245
5.8966 | -101.8237
5.9003 | -101.8186
5.904 | -101.8085
5.9077 | -101.8023
5.9114 | -101.7965
5.9151 | -101.7954
5.9188 | -101.7975
5.9225 | -101.8015
5.9262 | -101.8145
5.9299 | -101.8354
5.9336 | -101.8573
5.9373 | -101.8808
5.941 | -101.9005
5.9447 | -101.9203
5.9484 | -101.9381
5.9521 | -101.94
5.9558 | -101.9375
5.9595 | -101.918
5.9632 | -101.8864
5.9669 | -101.8489
5.9706 | -101.8028
5.9744 | -101.7468
5.9781 | -101.6866
5.9818 | -101.6282
5.9855 | -101.5734
5.9892 | -101.5301
5.9929 | -101.496
5.9966 | -101.4766
6.0003 | -101.4708
6.004 | -101.4801
6.0077 | -101.5084
6.0114 | -101.557
6.0151 | -101.6109
6.0188 | -101.669
6.0225 | -101.7333
6.0262 | -101.7951
6.0299 | -101.844
6.0336 | -101.8834
6.0373 | -101.9134
6.041 | -101.933
6.0447 | -101.9403
6.0484 | -101.9349
6.0521 | -101.9273
6.0558 | -101.9119
6.0595 | -101.8893
6.0632 | -101.8728
6.0669 | -101.8598
6.0707 | -101.849
6.0744 | -101.841
6.0781 | -101.8444
6.0818 | -101.854
6.0855 | -101.8685
6.0892 | -101.8833
6.0929 | -101.8941
6.0966 | -101.9008
6.1003 | -101.9062
6.104 | -101.9056
6.1077 | -101.9124
6.1114 | -101.9204
6.1151 | -101.9308
6.1188 | -101.9505
6.1225 | -101.9767
6.1262 | -102.0132
6.1299 | -102.0511
6.1336 | -102.0877
6.1373 | -102.124
6.141 | -102.1605
6.1447 | -102.1905
6.1484 | -102.2174
6.1521 | -102.2398
6.1558 | -102.2575
6.1595 | -102.2787
6.1632 | -102.2986
6.167 | -102.3045
6.1707 | -102.3062
6.1744 | -102.3116
6.1781 | -102.3208
6.1818 | -102.3288
6.1855 | -102.3398
6.1892 | -102.3562
6.1929 | -102.3686
6.1966 | -102.3867
6.2003 | -102.4071
6.204 | -102.4351
6.2077 | -102.4694
6.2114 | -102.5
6.2151 | -102.5343
6.2188 | -102.5697
6.2225 | -102.607
6.2262 | -102.6426
6.2299 | -102.6668
6.2336 | -102.6743
6.2373 | -102.6689
6.241 | -102.6551
6.2447 | -102.628
6.2484 | -102.59
6.2521 | -102.5548
6.2558 | -102.5286
6.2595 | -102.5082
6.2633 | -102.4824
6.267 | -102.4577
6.2707 | -102.4364
6.2744 | -102.4219
6.2781 | -102.4078
6.2818 | -102.3878
6.2855 | -102.3694
6.2892 | -102.3491
6.2929 | -102.3338
6.2966 | -102.3197
6.3003 | -102.3077
6.304 | -102.3087
6.3077 | -102.3103
6.3114 | -102.3156
6.3151 | -102.3291
6.3188 | -102.3405
6.3225 | -102.3483
6.3262 | -102.3535
6.3299 | -102.3517
6.3336 | -102.3392
6.3373 | -102.3224
6.341 | -102.3001
6.3447 | -102.2687
6.3484 | -102.2381
6.3521 | -102.2202
6.3559 | -102.2053
6.3596 | -102.1971
6.3633 | -102.1965
6.367 | -102.2015
6.3707 | -102.2183
6.3744 | -102.2454
6.3781 | -102.2753
6.3818 | -102.3071
6.3855 | -102.331
6.3892 | -102.3566
6.3929 | -102.3853
6.3966 | -102.4053
6.4003 | -102.4255
6.404 | -102.4372
6.4077 | -102.4421
6.4114 | -102.4337
6.4151 | -102.4284
6.4188 | -102.4179
6.4225 | -102.3996
6.4262 | -102.38
6.4299 | -102.3598
6.4336 | -102.3378
6.4373 | -102.3221
6.441 | -102.3162
6.4447 | -102.3213
6.4484 | -102.3301
6.4522 | -102.3473
6.4559 | -102.3807
6.4596 | -102.4226
6.4633 | -102.4769
6.467 | -102.5372
6.4707 | -102.5991
6.4744 | -102.6591
6.4781 | -102.7154
6.4818 | -102.7653
6.4855 | -102.8007
6.4892 | -102.8211
6.4929 | -102.8349
6.4966 | -102.8367
6.5003 | -102.8307
6.504 | -102.8262
6.5077 | -102.8181
6.5114 | -102.8144
6.5151 | -102.8195
6.5188 | -102.8267
6.5225 | -102.8298
6.5262 | -102.8241
6.5299 | -102.8115
6.5336 | -102.7967
6.5373 | -102.7715
6.541 | -102.7372
6.5447 | -102.6981
6.5485 | -102.6586
6.5522 | -102.6157
6.5559 | -102.5627
6.5596 | -102.5207
6.5633 | -102.4834
6.567 | -102.4532
6.5707 | -102.438
6.5744 | -102.4322
6.5781 | -102.4415
6.5818 | -102.4644
6.5855 | -102.4946
6.5892 | -102.536
6.5929 | -102.5855
6.5966 | -102.6368
6.6003 | -102.6835
6.604 | -102.7227
6.6077 | -102.7559
6.6114 | -102.7748
6.6151 | -102.7779
6.6188 | -102.767
6.6225 | -102.7487
6.6262 | -102.7265
6.6299 | -102.698
6.6336 | -102.6659
6.6373 | -102.6407
6.641 | -102.6256
6.6448 | -102.6245
6.6485 | -102.638
6.6522 | -102.6676
6.6559 | -102.7066
6.6596 | -102.7508
6.6633 | -102.8025
6.667 | -102.8571
6.6707 | -102.9126
6.6744 | -102.9698
6.6781 | -103.0157
6.6818 | -103.0498
6.6855 | -103.0773
6.6892 | -103.0947
6.6929 | -103.0995
6.6966 | -103.0972
6.7003 | -103.0815
6.704 | -103.0498
6.7077 | -103.0118
6.7114 | -102.9641
6.7151 | -102.9083
6.7188 | -102.8549
6.7225 | -102.8059
6.7262 | -102.7609
6.7299 | -102.7296
6.7336 | -102.7075
6.7373 | -102.6957
6.7411 | -102.7001
6.7448 | -102.7138
6.7485 | -102.7423
6.7522 | -102.7839
6.7559 | -102.838
6.7596 | -102.8958
6.7633 | -102.9517
6.767 | -103.0044
6.7707 | -103.0504
6.7744 | -103.0903
6.7781 | -103.12
6.7818 | -103.1434
6.7855 | -103.1449
6.7892 | -103.1221
6.7929 | -103.0853
6.7966 | -103.0399
6.8003 | -102.9917
6.804 | -102.9375
6.8077 | -102.8829
6.8114 | -102.8382
6.8151 | -102.802
6.8188 | -102.78
6.8225 | -102.7689
6.8262 | -102.7727
6.8299 | -102.7895
6.8336 | -102.8168
6.8374 | -102.8557
6.8411 | -102.8947
6.8448 | -102.9394
6.8485 | -102.989
6.8522 | -103.035
6.8559 | -103.0781
6.8596 | -103.1205
6.8633 | -103.1682
6.867 | -103.2209
6.8707 | -103.272
6.8744 | -103.3163
6.8781 | -103.3585
6.8818 | -103.3954
6.8855 | -103.4281
6.8892 | -103.458
6.8929 | -103.476
6.8966 | -103.485
6.9003 | -103.4886
6.904 | -103.4839
6.9077 | -103.472
6.9114 | -103.4549
6.9151 | -103.4336
6.9188 | -103.4055
6.9225 | -103.369
6.9262 | -103.3361
6.9299 | -103.3037
6.9337 | -103.274
6.9374 | -103.2481
6.9411 | -103.2267
6.9448 | -103.2242
6.9485 | -103.2146
6.9522 | -103.2058
6.9559 | -103.2054
6.9596 | -103.211
6.9633 | -103.2166
6.967 | -103.22
6.9707 | -103.2245
6.9744 | -103.229
6.9781 | -103.2341
6.9818 | -103.2382
6.9855 | -103.2406
6.9892 | -103.2451
6.9929 | -103.2444
6.9966 | -103.2378
7.0003 | -103.2214
7.004 | -103.2013
7.0077 | -103.1838
7.0114 | -103.1614
7.0151 | -103.1405
7.0188 | -103.1235
7.0225 | -103.1043
7.0263 | -103.0933
7.03 | -103.0846
7.0337 | -103.0804
7.0374 | -103.075
7.0411 | -103.0743
7.0448 | -103.0839
7.0485 | -103.0942
7.0522 | -103.1077
7.0559 | -103.1237
7.0596 | -103.1454
7.0633 | -103.1685
7.067 | -103.1923
7.0707 | -103.2177
7.0744 | -103.239
7.0781 | -103.2511
7.0818 | -103.2671
7.0855 | -103.28
7.0892 | -103.2882
7.0929 | -103.2953
7.0966 | -103.3016
7.1003 | -103.3092
7.104 | -103.3193
7.1077 | -103.333
7.1114 | -103.3442
7.1151 | -103.3459
7.1188 | -103.3472
7.1226 | -103.3455
7.1263 | -103.3352
7.13 | -103.3251
7.1337 | -103.3086
7.1374 | -103.2831
7.1411 | -103.255
7.1448 | -103.2323
7.1485 | -103.2056
7.1522 | -103.1698
7.1559 | -103.1374
7.1596 | -103.1097
7.1633 | -103.0849
7.167 | -103.0607
7.1707 | -103.0378
7.1744 | -103.0179
7.1781 | -103.0048
7.1818 | -103.001
7.1855 | -103.0072
7.1892 | -103.0221
7.1929 | -103.0494
7.1966 | -103.0832
7.2003 | -103.1244
7.204 | -103.1806
7.2077 | -103.238
7.2114 | -103.2984
7.2151 | -103.3618
7.2189 | -103.4258
7.2226 | -103.4868
7.2263 | -103.5383
7.23 | -103.5765
7.2337 | -103.6047
7.2374 | -103.622
7.2411 | -103.6305
7.2448 | -103.6286
7.2485 | -103.6153
7.2522 | -103.5897
7.2559 | -103.5547
7.2596 | -103.5201
7.2633 | -103.479
7.267 | -103.433
7.2707 | -103.39
7.2744 | -103.3493
7.2781 | -103.3115
7.2818 | -103.2785
7.2855 | -103.2451
7.2892 | -103.2174
7.2929 | -103.1866
7.2966 | -103.1575
7.3003 | -103.136
7.304 | -103.1231
7.3077 | -103.1195
7.3114 | -103.1148
7.3152 | -103.1245
7.3189 | -103.1452
7.3226 | -103.1682
7.3263 | -103.1927
7.33 | -103.227
7.3337 | -103.27
7.3374 | -103.3166
7.3411 | -103.3599
7.3448 | -103.4026
7.3485 | -103.4381
7.3522 | -103.4649
7.3559 | -103.4826
7.3596 | -103.4918
7.3633 | -103.4912
7.367 | -103.4805
7.3707 | -103.4528
7.3744 | -103.4105
7.3781 | -103.3645
7.3818 | -103.3123
7.3855 | -103.2578
7.3892 | -103.2022
7.3929 | -103.1444
7.3966 | -103.0853
7.4003 | -103.0394
7.404 | -102.9938
7.4078 | -102.9507
7.4115 | -102.9153
7.4152 | -102.8854
7.4189 | -102.8693
7.4226 | -102.8614
7.4263 | -102.8611
7.43 | -102.8728
7.4337 | -102.8989
7.4374 | -102.9304
7.4411 | -102.9694
7.4448 | -103.0177
7.4485 | -103.0762
7.4522 | -103.1409
7.4559 | -103.213
7.4596 | -103.2932
7.4633 | -103.3766
7.467 | -103.4598
7.4707 | -103.5438
7.4744 | -103.6253
7.4781 | -103.687
7.4818 | -103.7614
7.4855 | -103.8245
7.4892 | -103.8741
7.4929 | -103.9177
7.4966 | -103.9523
7.5003 | -103.9728
7.5041 | -103.9904
7.5078 | -104.0021
7.5115 | -104.0033
7.5152 | -104.001
7.5189 | -103.9952
7.5226 | -103.9868
7.5263 | -103.9801
7.53 | -103.9663
7.5337 | -103.9557
7.5374 | -103.9487
7.5411 | -103.9426
7.5448 | -103.9351
7.5485 | -103.9287
7.5522 | -103.92
7.5559 | -103.9134
7.5596 | -103.9042
7.5633 | -103.8931
7.567 | -103.8817
7.5707 | -103.8665
7.5744 | -103.8579
7.5781 | -103.8537
7.5818 | -103.8418
7.5855 | -103.8295
7.5892 | -103.8176
7.5929 | -103.8072
7.5966 | -103.7998
7.6004 | -103.7954
7.6041 | -103.7852
7.6078 | -103.7747
7.6115 | -103.7665
7.6152 | -103.7625
7.6189 | -103.7603
7.6226 | -103.7566
7.6263 | -103.7549
7.63 | -103.7612
7.6337 | -103.7713
7.6374 | -103.7786
7.6411 | -103.7859
7.6448 | -103.7995
7.6485 | -103.807
7.6522 | -103.8067
7.6559 | -103.8016
7.6596 | -103.7923
7.6633 | -103.7719
7.667 | -103.7486
7.6707 | -103.7241
7.6744 | -103.7035
7.6781 | -103.6769
7.6818 | -103.6465
7.6855 | -103.6158
7.6892 | -103.5786
7.6929 | -103.548
7.6967 | -103.5275
7.7004 | -103.5149
7.7041 | -103.507
7.7078 | -103.4992
7.7115 | -103.497
7.7152 | -103.5106
7.7189 | -103.5341
7.7226 | -103.5644
7.7263 | -103.5965
7.73 | -103.6351
7.7337 | -103.6802
7.7374 | -103.7279
7.7411 | -103.7768
7.7448 | -103.8214
7.7485 | -103.8653
7.7522 | -103.9021
7.7559 | -103.9288
7.7596 | -103.9491
7.7633 | -103.9636
7.767 | -103.9697
7.7707 | -103.9702
7.7744 | -103.9646
7.7781 | -103.9543
7.7818 | -103.9317
7.7855 | -103.9028
7.7893 | -103.869
7.793 | -103.8365
7.7967 | -103.8071
7.8004 | -103.7791
7.8041 | -103.7542
7.8078 | -103.7226
7.8115 | -103.6924
7.8152 | -103.6679
7.8189 | -103.6494
7.8226 | -103.6342
7.8263 | -103.6199
7.83 | -103.6029
7.8337 | -103.5885
7.8374 | -103.5757
7.8411 | -103.5714
7.8448 | -103.5699
7.8485 | -103.5696
7.8522 | -103.5731
7.8559 | -103.5797
7.8596 | -103.5886
7.8633 | -103.5998
7.867 | -103.6164
7.8707 | -103.6376
7.8744 | -103.6625
7.8781 | -103.6911
7.8818 | -103.7221
7.8856 | -103.7527
7.8893 | -103.781
7.893 | -103.8124
7.8967 | -103.8471
7.9004 | -103.8733
7.9041 | -103.9027
7.9078 | -103.9393
7.9115 | -103.9791
7.9152 | -104.0177
7.9189 | -104.0582
7.9226 | -104.1002
7.9263 | -104.145
7.93 | -104.1844
7.9337 | -104.2252
7.9374 | -104.2687
7.9411 | -104.3083
7.9448 | -104.3428
7.9485 | -104.3709
7.9522 | -104.3906
7.9559 | -104.3963
7.9596 | -104.3898
7.9633 | -104.375
7.967 | -104.3586
7.9707 | -104.3388
7.9744 | -104.318
7.9781 | -104.2916
7.9819 | -104.2669
7.9856 | -104.239
7.9893 | -104.2105
7.993 | -104.1896
7.9967 | -104.175
8.0004 | -104.1641
8.0041 | -104.1589
8.0078 | -104.1569
8.0115 | -104.16
8.0152 | -104.1731
8.0189 | -104.1865
8.0226 | -104.1948
8.0263 | -104.1981
8.03 | -104.2026
8.0337 | -104.2074
8.0374 | -104.2172
8.0411 | -104.2246
8.0448 | -104.2223
8.0485 | -104.2202
8.0522 | -104.2178
8.0559 | -104.2149
8.0596 | -104.2096
8.0633 | -104.2031
8.067 | -104.1995
8.0707 | -104.1994
8.0744 | -104.1991
8.0782 | -104.1968
8.0819 | -104.1944
8.0856 | -104.1916
8.0893 | -104.1904
8.093 | -104.1856
8.0967 | -104.1863
8.1004 | -104.1886
8.1041 | -104.1942
8.1078 | -104.2053
8.1115 | -104.2171
8.1152 | -104.2358
8.1189 | -104.2609
8.1226 | -104.2884
8.1263 | -104.3117
8.13 | -104.3328
8.1337 | -104.3565
8.1374 | -104.3757
8.1411 | -104.3886
8.1448 | -104.398
8.1485 | -104.3998
8.1522 | -104.3924
8.1559 | -104.3787
8.1596 | -104.3571
8.1633 | -104.3361
8.167 | -104.3105
8.1707 | -104.2798
8.1745 | -104.249
8.1782 | -104.2198
8.1819 | -104.1983
8.1856 | -104.1833
8.1893 | -104.1732
8.193 | -104.1643
8.1967 | -104.1632
8.2004 | -104.1657
8.2041 | -104.1665
8.2078 | -104.1674
8.2115 | -104.1688
8.2152 | -104.1755
8.2189 | -104.1755
8.2226 | -104.1744
8.2263 | -104.1715
8.23 | -104.1643
8.2337 | -104.1583
8.2374 | -104.15
8.2411 | -104.1387
8.2448 | -104.1263
8.2485 | -104.1207
8.2522 | -104.1159
8.2559 | -104.1041
8.2596 | -104.0932
8.2633 | -104.0877
8.267 | -104.0798
8.2708 | -104.0755
8.2745 | -104.0753
8.2782 | -104.0761
8.2819 | -104.077
8.2856 | -104.0774
8.2893 | -104.0816
8.293 | -104.0856
8.2967 | -104.085
8.3004 | -104.0879
8.3041 | -104.0918
8.3078 | -104.0959
8.3115 | -104.1015
8.3152 | -104.1118
8.3189 | -104.1246
8.3226 | -104.1425
8.3263 | -104.1643
8.33 | -104.1842
8.3337 | -104.2057
8.3374 | -104.2268
8.3411 | -104.2517
8.3448 | -104.2701
8.3485 | -104.2862
8.3522 | -104.3009
8.3559 | -104.3092
8.3596 | -104.3153
8.3634 | -104.3268
8.3671 | -104.3315
8.3708 | -104.3305
8.3745 | -104.3252
8.3782 | -104.3165
8.3819 | -104.3072
8.3856 | -104.2912
8.3893 | -104.2724
8.393 | -104.2526
8.3967 | -104.2311
8.4004 | -104.2064
8.4041 | -104.1819
8.4078 | -104.1633
8.4115 | -104.1455
8.4152 | -104.1291
8.4189 | -104.1175
8.4226 | -104.1057
8.4263 | -104.0982
8.43 | -104.0952
8.4337 | -104.0966
8.4374 | -104.1047
8.4411 | -104.1176
8.4448 | -104.1349
8.4485 | -104.1568
8.4522 | -104.1842
8.4559 | -104.2109
8.4597 | -104.2386
8.4634 | -104.2661
8.4671 | -104.2968
8.4708 | -104.3305
8.4745 | -104.3641
8.4782 | -104.3959
8.4819 | -104.4199
8.4856 | -104.4348
8.4893 | -104.4591
8.493 | -104.4802
8.4967 | -104.4934
8.5004 | -104.5092
8.5041 | -104.5164
8.5078 | -104.5165
8.5115 | -104.517
8.5152 | -104.5121
8.5189 | -104.4995
8.5226 | -104.4894
8.5263 | -104.4774
8.53 | -104.4619
8.5337 | -104.4515
8.5374 | -104.4429
8.5411 | -104.4357
8.5448 | -104.4267
8.5485 | -104.4179
8.5522 | -104.4098
8.556 | -104.3985
8.5597 | -104.3855
8.5634 | -104.373
8.5671 | -104.3598
8.5708 | -104.3471
8.5745 | -104.3306
8.5782 | -104.3177
8.5819 | -104.3078
8.5856 | -104.2973
8.5893 | -104.2931
8.593 | -104.2887
8.5967 | -104.2822
8.6004 | -104.2854
8.6041 | -104.2941
8.6078 | -104.3079
8.6115 | -104.3258
8.6152 | -104.3362
8.6189 | -104.3456
8.6226 | -104.3617
8.6263 | -104.3707
8.63 | -104.3772
8.6337 | -104.3819
8.6374 | -104.3807
8.6411 | -104.3704
8.6448 | -104.3551
8.6485 | -104.3336
8.6523 | -104.3122
8.656 | -104.2883
8.6597 | -104.2615
8.6634 | -104.2375
8.6671 | -104.2162
8.6708 | -104.1975
8.6745 | -104.182
8.6782 | -104.1656
8.6819 | -104.1568
8.6856 | -104.1563
8.6893 | -104.1608
8.693 | -104.1673
8.6967 | -104.1757
8.7004 | -104.1856
8.7041 | -104.202
8.7078 | -104.2186
8.7115 | -104.2406
8.7152 | -104.2682
8.7189 | -104.2842
8.7226 | -104.3006
8.7263 | -104.3217
8.73 | -104.3308
8.7337 | -104.3389
8.7374 | -104.3513
8.7411 | -104.3605
8.7448 | -104.3704
8.7486 | -104.3825
8.7523 | -104.3911
8.756 | -104.4003
8.7597 | -104.4098
8.7634 | -104.4214
8.7671 | -104.4331
8.7708 | -104.4521
8.7745 | -104.4746
8.7782 | -104.4974
8.7819 | -104.5136
8.7856 | -104.5318
8.7893 | -104.5477
8.793 | -104.5595
8.7967 | -104.5666
8.8004 | -104.5693
8.8041 | -104.5669
8.8078 | -104.5671
8.8115 | -104.5605
8.8152 | -104.5494
8.8189 | -104.5385
8.8226 | -104.523
8.8263 | -104.5064
8.83 | -104.4902
8.8337 | -104.4722
8.8374 | -104.4523
8.8412 | -104.4297
8.8449 | -104.4119
8.8486 | -104.404
8.8523 | -104.397
8.856 | -104.3855
8.8597 | -104.3803
8.8634 | -104.3818
8.8671 | -104.3808
8.8708 | -104.3803
8.8745 | -104.3837
8.8782 | -104.3933
8.8819 | -104.4058
8.8856 | -104.4255
8.8893 | -104.447
8.893 | -104.4668
8.8967 | -104.4869
8.9004 | -104.5034
8.9041 | -104.5202
8.9078 | -104.5368
8.9115 | -104.5521
8.9152 | -104.5668
8.9189 | -104.5774
8.9226 | -104.5866
8.9263 | -104.5993
8.93 | -104.6134
8.9337 | -104.6255
8.9375 | -104.6301
8.9412 | -104.6349
8.9449 | -104.6414
8.9486 | -104.6459
8.9523 | -104.6514
8.956 | -104.6483
8.9597 | -104.6434
8.9634 | -104.6366
8.9671 | -104.6246
8.9708 | -104.6091
8.9745 | -104.584
8.9782 | -104.558
8.9819 | -104.5299
8.9856 | -104.5005
8.9893 | -104.4671
8.993 | -104.429
8.9967 | -104.397
9.0004 | -104.3702
9.0041 | -104.3441
9.0078 | -104.3136
9.0115 | -104.2877
9.0152 | -104.2632
9.0189 | -104.2439
9.0226 | -104.2286
9.0263 | -104.2188
9.03 | -104.2154
9.0338 | -104.219
9.0375 | -104.2284
9.0412 | -104.2448
9.0449 | -104.2699
9.0486 | -104.3055
9.0523 | -104.3476
9.056 | -104.3897
9.0597 | -104.4329
9.0634 | -104.4778
9.0671 | -104.5271
9.0708 | -104.5747
9.0745 | -104.6183
9.0782 | -104.6584
9.0819 | -104.6719
9.0856 | -104.7003
9.0893 | -104.7284
9.093 | -104.7485
9.0967 | -104.7619
9.1004 | -104.7732
9.1041 | -104.7779
9.1078 | -104.7806
9.1115 | -104.7795
9.1152 | -104.7813
9.1189 | -104.7825
9.1226 | -104.7842
9.1263 | -104.7853
9.1301 | -104.7878
9.1338 | -104.7898
9.1375 | -104.7942
9.1412 | -104.8032
9.1449 | -104.8102
9.1486 | -104.8174
9.1523 | -104.828
9.156 | -104.8383
9.1597 | -104.8459
9.1634 | -104.8513
9.1671 | -104.8564
9.1708 | -104.8606
9.1745 | -104.8629
9.1782 | -104.8582
9.1819 | -104.8533
9.1856 | -104.8478
9.1893 | -104.8454
9.193 | -104.8421
9.1967 | -104.8387
9.2004 | -104.8304
9.2041 | -104.8214
9.2078 | -104.8123
9.2115 | -104.803
9.2152 | -104.7973
9.2189 | -104.7875
9.2226 | -104.779
9.2264 | -104.7725
9.2301 | -104.7658
9.2338 | -104.7606
9.2375 | -104.7569
9.2412 | -104.7584
9.2449 | -104.7573
9.2486 | -104.754
9.2523 | -104.7527
9.256 | -104.7518
9.2597 | -104.7509
9.2634 | -104.7491
9.2671 | -104.7408
9.2708 | -104.7355
9.2745 | -104.7287
9.2782 | -104.7181
9.2819 | -104.708
9.2856 | -104.6939
9.2893 | -104.6746
9.293 | -104.6572
9.2967 | -104.638
9.3004 | -104.6147
9.3041 | -104.5867
9.3078 | -104.5601
9.3115 | -104.5323
9.3152 | -104.5061
9.3189 | -104.4832
9.3227 | -104.4599
9.3264 | -104.4435
9.3301 | -104.434
9.3338 | -104.425
9.3375 | -104.4242
9.3412 | -104.4292
9.3449 | -104.4385
9.3486 | -104.4537
9.3523 | -104.4738
9.356 | -104.4943
9.3597 | -104.5126
9.3634 | -104.5379
9.3671 | -104.5639
9.3708 | -104.5906
9.3745 | -104.6168
9.3782 | -104.6405
9.3819 | -104.6657
9.3856 | -104.6895
9.3893 | -104.7065
9.393 | -104.7238
9.3967 | -104.7401
9.4004 | -104.7465
9.4041 | -104.7504
9.4078 | -104.75
9.4115 | -104.7485
9.4153 | -104.7444
9.419 | -104.7385
9.4227 | -104.7321
9.4264 | -104.7218
9.4301 | -104.7082
9.4338 | -104.6961
9.4375 | -104.6875
9.4412 | -104.68
9.4449 | -104.6726
9.4486 | -104.6679
9.4523 | -104.6651
9.456 | -104.6576
9.4597 | -104.6526
9.4634 | -104.6461
9.4671 | -104.6398
9.4708 | -104.6358
9.4745 | -104.6361
9.4782 | -104.6368
9.4819 | -104.642
9.4856 | -104.65
9.4893 | -104.659
9.493 | -104.6711
9.4967 | -104.6824
9.5004 | -104.6885
9.5041 | -104.6934
9.5078 | -104.6959
9.5116 | -104.6967
9.5153 | -104.699
9.519 | -104.6972
9.5227 | -104.6935
9.5264 | -104.6891
9.5301 | -104.6856
9.5338 | -104.6831
9.5375 | -104.6846
9.5412 | -104.6878
9.5449 | -104.6917
9.5486 | -104.6983
9.5523 | -104.7133
9.556 | -104.7258
9.5597 | -104.743
9.5634 | -104.7626
9.5671 | -104.7843
9.5708 | -104.8114
9.5745 | -104.8424
9.5782 | -104.8768
9.5819 | -104.9124
9.5856 | -104.9496
9.5893 | -104.9825
9.593 | -105.0193
9.5967 | -105.0589
9.6004 | -105.0964
9.6041 | -105.1323
9.6079 | -105.1651
9.6116 | -105.1901
9.6153 | -105.1929
9.619 | -105.2059
9.6227 | -105.2154
9.6264 | -105.219
9.6301 | -105.2222
9.6338 | -105.2242
9.6375 | -105.2247
9.6412 | -105.2216
9.6449 | -105.2173
9.6486 | -105.2132
9.6523 | -105.2074
9.656 | -105.1987
9.6597 | -105.1925
9.6634 | -105.1785
9.6671 | -105.1639
9.6708 | -105.1486
9.6745 | -105.1357
9.6782 | -105.1273
9.6819 | -105.117
9.6856 | -105.1057
9.6893 | -105.0947
9.693 | -105.082
9.6967 | -105.0726
9.7004 | -105.0635
9.7042 | -105.0598
9.7079 | -105.0573
9.7116 | -105.0499
9.7153 | -105.0381
9.719 | -105.0255
9.7227 | -105.0173
9.7264 | -105.0122
9.7301 | -105.0073
9.7338 | -105.0035
9.7375 | -104.9987
9.7412 | -104.9925
9.7449 | -104.9867
9.7486 | -104.9836
9.7523 | -104.9823
9.756 | -104.9828
9.7597 | -104.9886
9.7634 | -104.9902
9.7671 | -104.9876
9.7708 | -104.985
9.7745 | -104.9831
9.7782 | -104.9845
9.7819 | -104.9859
9.7856 | -104.9893
9.7893 | -104.9948
9.793 | -105.0008
9.7967 | -105.0095
9.8005 | -105.0201
9.8042 | -105.0347
9.8079 | -105.0566
9.8116 | -105.076
9.8153 | -105.0946
9.819 | -105.1159
9.8227 | -105.1402
9.8264 | -105.1628
9.8301 | -105.1847
9.8338 | -105.2066
9.8375 | -105.2254
9.8412 | -105.247
9.8449 | -105.2686
9.8486 | -105.2846
9.8523 | -105.2982
9.856 | -105.3151
9.8597 | -105.3322
9.8634 | -105.3469
9.8671 | -105.3549
9.8708 | -105.3579
9.8745 | -105.3644
9.8782 | -105.3709
9.8819 | -105.3773
9.8856 | -105.3837
9.8893 | -105.3942
9.8931 | -105.4034
9.8968 | -105.4064
9.9005 | -105.4122
9.9042 | -105.4181
9.9079 | -105.4254
9.9116 | -105.4335
9.9153 | -105.4469
9.919 | -105.4569
9.9227 | -105.4627
9.9264 | -105.4693
9.9301 | -105.4777
9.9338 | -105.4846
9.9375 | -105.488
9.9412 | -105.4881
9.9449 | -105.4859
9.9486 | -105.4815
9.9523 | -105.4752
9.956 | -105.4693
9.9597 | -105.4579
9.9634 | -105.4349
9.9671 | -105.4116
9.9708 | -105.3904
9.9745 | -105.3646
9.9782 | -105.3355
9.9819 | -105.307
9.9856 | -105.2792
9.9894 | -105.2529
9.9931 | -105.23
9.9968 | -105.2072
10.0005 | -105.1849
10.0042 | -105.1721
10.0079 | -105.1633
10.0116 | -105.1584
10.0153 | -105.1612
10.019 | -105.1672
10.0227 | -105.1717
10.0264 | -105.1782
10.0301 | -105.19
10.0338 | -105.2003
10.0375 | -105.2153
10.0412 | -105.2365
10.0449 | -105.2562
10.0486 | -105.2734
10.0523 | -105.2883
10.056 | -105.3012
10.0597 | -105.3106
10.0634 | -105.3184
10.0671 | -105.3264
10.0708 | -105.33
10.0745 | -105.3266
10.0782 | -105.3223
10.0819 | -105.3137
10.0857 | -105.3013
10.0894 | -105.285
10.0931 | -105.2618
10.0968 | -105.2357
10.1005 | -105.206
10.1042 | -105.1763
10.1079 | -105.1511
10.1116 | -105.1204
10.1153 | -105.0892
10.119 | -105.0571
10.1227 | -105.0248
10.1264 | -104.9979
10.1301 | -104.9703
10.1338 | -104.9423
10.1375 | -104.9194
10.1412 | -104.9068
10.1449 | -104.8982
10.1486 | -104.8884
10.1523 | -104.9058
10.156 | -104.9091
10.1597 | -104.9177
10.1634 | -104.9284
10.1671 | -104.9444
10.1708 | -104.9666
10.1745 | -104.9932
10.1782 | -105.0247
10.182 | -105.06
10.1857 | -105.0932
10.1894 | -105.1272
10.1931 | -105.159
10.1968 | -105.1886
10.2005 | -105.2167
10.2042 | -105.2432
10.2079 | -105.2632
10.2116 | -105.2792
10.2153 | -105.2904
10.219 | -105.2969
10.2227 | -105.3031
10.2264 | -105.3058
10.2301 | -105.306
10.2338 | -105.3044
10.2375 | -105.297
10.2412 | -105.2894
10.2449 | -105.2808
10.2486 | -105.2716
10.2523 | -105.2625
10.256 | -105.2542
10.2597 | -105.2455
10.2634 | -105.2402
10.2671 | -105.2447
10.2708 | -105.2478
10.2745 | -105.2511
10.2783 | -105.2586
10.282 | -105.2692
10.2857 | -105.284
10.2894 | -105.2981
10.2931 | -105.3137
10.2968 | -105.3322
10.3005 | -105.3508
10.3042 | -105.3675
10.3079 | -105.3836
10.3116 | -105.4019
10.3153 | -105.4199
10.319 | -105.4329
10.3227 | -105.4444
10.3264 | -105.4567
10.3301 | -105.4658
10.3338 | -105.4719
10.3375 | -105.475
10.3412 | -105.4751
10.3449 | -105.4767
10.3486 | -105.4765
10.3523 | -105.4736
10.356 | -105.4657
10.3597 | -105.4557
10.3634 | -105.4431
10.3671 | -105.4271
10.3709 | -105.412
10.3746 | -105.3911
10.3783 | -105.372
10.382 | -105.3543
10.3857 | -105.3331
10.3894 | -105.3143
10.3931 | -105.294
10.3968 | -105.2731
10.4005 | -105.2551
10.4042 | -105.241
10.4079 | -105.2295
10.4116 | -105.2214
10.4153 | -105.2143
10.419 | -105.206
10.4227 | -105.2026
10.4264 | -105.2017
10.4301 | -105.2054
10.4338 | -105.2142
10.4375 | -105.2214
10.4412 | -105.2327
10.4449 | -105.2499
10.4486 | -105.2712
10.4523 | -105.2933
10.456 | -105.3134
10.4597 | -105.3346
10.4634 | -105.3537
10.4672 | -105.3742
10.4709 | -105.395
10.4746 | -105.4125
10.4783 | -105.423
10.482 | -105.4301
10.4857 | -105.4369
10.4894 | -105.4434
10.4931 | -105.4454
10.4968 | -105.4454
10.5005 | -105.4449
10.5042 | -105.4388
10.5079 | -105.4317
10.5116 | -105.4273
10.5153 | -105.4241
10.519 | -105.4238
10.5227 | -105.4212
10.5264 | -105.4208
10.5301 | -105.4165
10.5338 | -105.4155
10.5375 | -105.421
10.5412 | -105.4283
10.5449 | -105.4365
10.5486 | -105.4434
10.5523 | -105.447
10.556 | -105.4509
10.5597 | -105.4556
10.5635 | -105.4644
10.5672 | -105.473
10.5709 | -105.4792
10.5746 | -105.4847
10.5783 | -105.4867
10.582 | -105.4879
10.5857 | -105.485
10.5894 | -105.4851
10.5931 | -105.4814
10.5968 | -105.4752
10.6005 | -105.466
10.6042 | -105.4593
10.6079 | -105.4582
10.6116 | -105.4546
10.6153 | -105.4523
10.619 | -105.4515
10.6227 | -105.4545
10.6264 | -105.4594
10.6301 | -105.4666
10.6338 | -105.4725
10.6375 | -105.4791
10.6412 | -105.488
10.6449 | -105.4974
10.6486 | -105.5055
10.6523 | -105.5124
10.656 | -105.523
10.6598 | -105.5305
10.6635 | -105.5326
10.6672 | -105.5347
10.6709 | -105.5336
10.6746 | -105.5303
10.6783 | -105.5252
10.682 | -105.5213
10.6857 | -105.5082
10.6894 | -105.5017
10.6931 | -105.4925
10.6968 | -105.4826
10.7005 | -105.476
10.7042 | -105.4735
10.7079 | -105.4709
10.7116 | -105.467
10.7153 | -105.4666
10.719 | -105.4695
10.7227 | -105.4798
10.7264 | -105.4947
10.7301 | -105.5103
10.7338 | -105.528
10.7375 | -105.5509
10.7412 | -105.5794
10.7449 | -105.6055
10.7486 | -105.6303
10.7523 | -105.654
10.7561 | -105.6797
10.7598 | -105.7016
10.7635 | -105.7238
10.7672 | -105.7476
10.7709 | -105.7644
10.7746 | -105.7779
10.7783 | -105.7872
10.782 | -105.7932
10.7857 | -105.8004
10.7894 | -105.8094
10.7931 | -105.8155
10.7968 | -105.8181
10.8005 | -105.8173
10.8042 | -105.8183
10.8079 | -105.8189
10.8116 | -105.8127
10.8153 | -105.8096
10.819 | -105.8076
10.8227 | -105.8023
10.8264 | -105.7941
10.8301 | -105.7866
10.8338 | -105.7729
10.8375 | -105.763
10.8412 | -105.7506
10.8449 | -105.7386
10.8486 | -105.7237
10.8524 | -105.7119
10.8561 | -105.7023
10.8598 | -105.6913
10.8635 | -105.6782
10.8672 | -105.667
10.8709 | -105.6564
10.8746 | -105.644
10.8783 | -105.6312
10.882 | -105.6161
10.8857 | -105.6029
10.8894 | -105.591
10.8931 | -105.5789
10.8968 | -105.5669
10.9005 | -105.5558
10.9042 | -105.549
10.9079 | -105.5473
10.9116 | -105.5432
10.9153 | -105.5387
10.919 | -105.5346
10.9227 | -105.5284
10.9264 | -105.5208
10.9301 | -105.5177
10.9338 | -105.5165
10.9375 | -105.5163
10.9412 | -105.5184
10.945 | -105.5201
10.9487 | -105.5191
10.9524 | -105.5154
10.9561 | -105.5132
10.9598 | -105.5119
10.9635 | -105.5098
10.9672 | -105.5123
10.9709 | -105.5213
10.9746 | -105.5281
10.9783 | -105.5356
10.982 | -105.549
10.9857 | -105.5649
10.9894 | -105.5807
10.9931 | -105.6025
10.9968 | -105.6198
11.0005 | -105.6397
11.0042 | -105.6591
11.0079 | -105.6848
11.0116 | -105.7083
11.0153 | -105.7343
11.019 | -105.7607
11.0227 | -105.7888
11.0264 | -105.8175
11.0301 | -105.8435
11.0338 | -105.8704
11.0375 | -105.8962
11.0413 | -105.9231
11.045 | -105.9462
11.0487 | -105.9645
11.0524 | -105.9826
11.0561 | -105.9933
11.0598 | -106.0004
11.0635 | -106.0027
11.0672 | -106.0015
11.0709 | -106.0006
11.0746 | -105.9935
11.0783 | -105.9807
11.082 | -105.9687
11.0857 | -105.9565
11.0894 | -105.9415
11.0931 | -105.9215
11.0968 | -105.9008
11.1005 | -105.8766
11.1042 | -105.8521
11.1079 | -105.8268
11.1116 | -105.8017
11.1153 | -105.7773
11.119 | -105.7547
11.1227 | -105.726
11.1264 | -105.6952
11.1301 | -105.6643
11.1338 | -105.6324
11.1376 | -105.6
11.1413 | -105.5702
11.145 | -105.5457
11.1487 | -105.5212
11.1524 | -105.4971
11.1561 | -105.4735
11.1598 | -105.4538
11.1635 | -105.4363
11.1672 | -105.4154
11.1709 | -105.3955
11.1746 | -105.3765
11.1783 | -105.3654
11.182 | -105.3617
11.1857 | -105.3604
11.1894 | -105.3612
11.1931 | -105.3605
11.1968 | -105.3612
11.2005 | -105.3635
11.2042 | -105.3677
11.2079 | -105.3735
11.2116 | -105.3804
11.2153 | -105.3893
11.219 | -105.3962
11.2227 | -105.3996
11.2264 | -105.4043
11.2302 | -105.4079
11.2339 | -105.4123
11.2376 | -105.4208
11.2413 | -105.433
11.245 | -105.439
11.2487 | -105.4455
11.2524 | -105.4494
11.2561 | -105.4572
11.2598 | -105.4652
11.2635 | -105.4725
11.2672 | -105.4783
11.2709 | -105.4911
11.2746 | -105.5014
11.2783 | -105.5059
11.282 | -105.5109
11.2857 | -105.5174
11.2894 | -105.5217
11.2931 | -105.5246
11.2968 | -105.5259
11.3005 | -105.5293
11.3042 | -105.5272
11.3079 | -105.5235
11.3116 | -105.5244
11.3153 | -105.526
11.319 | -105.5305
11.3227 | -105.536
11.3264 | -105.5425
11.3302 | -105.5528
11.3339 | -105.5665
11.3376 | -105.5828
11.3413 | -105.5943
11.345 | -105.6004
11.3487 | -105.6132
11.3524 | -105.6287
11.3561 | -105.6455
11.3598 | -105.6634
11.3635 | -105.6812
11.3672 | -105.7
11.3709 | -105.7177
11.3746 | -105.7307
11.3783 | -105.7407
11.382 | -105.7499
11.3857 | -105.7562
11.3894 | -105.7622
11.3931 | -105.7636
11.3968 | -105.7623
11.4005 | -105.7599
11.4042 | -105.7528
11.4079 | -105.7399
11.4116 | -105.7251
11.4153 | -105.7138
11.419 | -105.7033
11.4228 | -105.6908
11.4265 | -105.6792
11.4302 | -105.6709
11.4339 | -105.6633
11.4376 | -105.6526
11.4413 | -105.6477
11.445 | -105.6481
11.4487 | -105.6431
11.4524 | -105.6358
11.4561 | -105.6276
11.4598 | -105.6233
11.4635 | -105.6201
11.4672 | -105.6149
11.4709 | -105.6085
11.4746 | -105.6015
11.4783 | -105.5943
11.482 | -105.5853
11.4857 | -105.5774
11.4894 | -105.5675
11.4931 | -105.5583
11.4968 | -105.5485
11.5005 | -105.5396
11.5042 | -105.5299
11.5079 | -105.5251
11.5116 | -105.5233
11.5153 | -105.5153
11.5191 | -105.5096
11.5228 | -105.5086
11.5265 | -105.5081
11.5302 | -105.5072
11.5339 | -105.5069
11.5376 | -105.5071
11.5413 | -105.5081
11.545 | -105.5138
11.5487 | -105.5203
11.5524 | -105.5292
11.5561 | -105.5394
11.5598 | -105.5496
11.5635 | -105.5582
11.5672 | -105.5623
11.5709 | -105.5688
11.5746 | -105.5752
11.5783 | -105.5783
11.582 | -105.5771
11.5857 | -105.5763
11.5894 | -105.5739
11.5931 | -105.5672
11.5968 | -105.5565
11.6005 | -105.5463
11.6042 | -105.5347
11.6079 | -105.5215
11.6116 | -105.506
11.6154 | -105.4946
11.6191 | -105.4867
11.6228 | -105.4792
11.6265 | -105.4706
11.6302 | -105.463
11.6339 | -105.4561
11.6376 | -105.4537
11.6413 | -105.4553
11.645 | -105.4565
11.6487 | -105.4614
11.6524 | -105.469
11.6561 | -105.4811
11.6598 | -105.4943
11.6635 | -105.5105
11.6672 | -105.5268
11.6709 | -105.5448
11.6746 | -105.5639
11.6783 | -105.5847
11.682 | -105.6091
11.6857 | -105.6333
11.6894 | -105.6568
11.6931 | -105.6819
11.6968 | -105.7076
11.7005 | -105.732
11.7042 | -105.7587
11.7079 | -105.7848
11.7117 | -105.8059
11.7154 | -105.828
11.7191 | -105.8457
11.7228 | -105.8584
11.7265 | -105.8706
11.7302 | -105.882
11.7339 | -105.8937
11.7376 | -105.9034
11.7413 | -105.9101
11.745 | -105.9178
11.7487 | -105.9189
11.7524 | -105.9095
11.7561 | -105.9068
11.7598 | -105.9026
11.7635 | -105.8999
11.7672 | -105.8962
11.7709 | -105.8919
11.7746 | -105.8868
11.7783 | -105.8826
11.782 | -105.8814
11.7857 | -105.8794
11.7894 | -105.8791
11.7931 | -105.8798
11.7968 | -105.8781
11.8005 | -105.8774
11.8042 | -105.8773
11.808 | -105.8796
11.8117 | -105.8829
11.8154 | -105.8867
11.8191 | -105.8899
11.8228 | -105.8908
11.8265 | -105.8902
11.8302 | -105.892
11.8339 | -105.8928
11.8376 | -105.8886
11.8413 | -105.8805
11.845 | -105.8685
11.8487 | -105.8556
11.8524 | -105.8391
11.8561 | -105.8171
11.8598 | -105.7905
11.8635 | -105.7627
11.8672 | -105.7354
11.8709 | -105.7102
11.8746 | -105.6819
11.8783 | -105.6524
11.882 | -105.6269
11.8857 | -105.6015
11.8894 | -105.579
11.8931 | -105.5581
11.8968 | -105.5364
11.9006 | -105.5161
11.9043 | -105.4974
11.908 | -105.4789
11.9117 | -105.4665
11.9154 | -105.4573
11.9191 | -105.4508
11.9228 | -105.4497
11.9265 | -105.4502
11.9302 | -105.4477
11.9339 | -105.4486
11.9376 | -105.4542
11.9413 | -105.4635
11.945 | -105.4737
11.9487 | -105.4896
11.9524 | -105.5024
11.9561 | -105.5138
11.9598 | -105.5305
11.9635 | -105.5506
11.9672 | -105.5739
11.9709 | -105.6004
11.9746 | -105.6256
11.9783 | -105.6507
11.982 | -105.6718
11.9857 | -105.6866
11.9894 | -105.7045
11.9931 | -105.7213
11.9969 | -105.7348
12.0006 | -105.7467
12.0043 | -105.7574
12.008 | -105.7636
12.0117 | -105.7722
12.0154 | -105.7796
12.0191 | -105.7818
12.0228 | -105.7798
12.0265 | -105.7769
12.0302 | -105.772
12.0339 | -105.7671
12.0376 | -105.7571
12.0413 | -105.7456
12.045 | -105.7339
12.0487 | -105.7203
12.0524 | -105.7072
12.0561 | -105.694
12.0598 | -105.6794
12.0635 | -105.6636
12.0672 | -105.6507
12.0709 | -105.6382
12.0746 | -105.6251
12.0783 | -105.6157
12.082 | -105.6086
12.0857 | -105.5987
12.0894 | -105.588
12.0932 | -105.5795
12.0969 | -105.5722
12.1006 | -105.5686
12.1043 | -105.5723
12.108 | -105.5749
12.1117 | -105.5774
12.1154 | -105.5838
12.1191 | -105.5946
12.1228 | -105.6035
12.1265 | -105.6198
12.1302 | -105.6389
12.1339 | -105.6542
12.1376 | -105.6683
12.1413 | -105.6888
12.145 | -105.7111
12.1487 | -105.7373
12.1524 | -105.7603
12.1561 | -105.785
12.1598 | -105.8107
12.1635 | -105.8324
12.1672 | -105.8566
12.1709 | -105.8824
12.1746 | -105.9055
12.1783 | -105.9208
12.182 | -105.9387
12.1857 | -105.954
12.1895 | -105.965
12.1932 | -105.9804
12.1969 | -105.9987
12.2006 | -106.0132
12.2043 | -106.0269
12.208 | -106.0359
12.2117 | -106.0433
12.2154 | -106.0484
12.2191 | -106.051
12.2228 | -106.0553
12.2265 | -106.0552
12.2302 | -106.0513
12.2339 | -106.049
12.2376 | -106.0472
12.2413 | -106.0411
12.245 | -106.0388
12.2487 | -106.0361
12.2524 | -106.0278
12.2561 | -106.0193
12.2598 | -106.0182
12.2635 | -106.0159
12.2672 | -106.0131
12.2709 | -106.0122
12.2746 | -106.0113
12.2783 | -106.0053
12.282 | -106.0041
12.2858 | -106.0044
12.2895 | -106.0013
12.2932 | -106.0004
12.2969 | -105.998
12.3006 | -105.9968
12.3043 | -105.9921
12.308 | -105.9906
12.3117 | -105.9919
12.3154 | -105.9916
12.3191 | -105.9869
12.3228 | -105.9829
12.3265 | -105.9826
12.3302 | -105.9883
12.3339 | -105.9926
12.3376 | -105.9973
12.3413 | -106.0023
12.345 | -106.0041
12.3487 | -106.0058
12.3524 | -106.0109
12.3561 | -106.0161
12.3598 | -106.017
12.3635 | -106.0194
12.3672 | -106.0253
12.3709 | -106.0276
12.3746 | -106.0258
12.3784 | -106.0246
12.3821 | -106.0229
12.3858 | -106.0222
12.3895 | -106.0215
12.3932 | -106.0178
12.3969 | -106.0121
12.4006 | -106.0053
12.4043 | -106.0005
12.408 | -105.9964
12.4117 | -105.9919
12.4154 | -105.9856
12.4191 | -105.9794
12.4228 | -105.9707
12.4265 | -105.9628
12.4302 | -105.9583
12.4339 | -105.9516
12.4376 | -105.9409
12.4413 | -105.9342
12.445 | -105.93
12.4487 | -105.9253
12.4524 | -105.9225
12.4561 | -105.9256
12.4598 | -105.9276
12.4635 | -105.9267
12.4672 | -105.9304
12.4709 | -105.9374
12.4747 | -105.9463
12.4784 | -105.9544
12.4821 | -105.967
12.4858 | -105.9866
12.4895 | -106.0041
12.4932 | -106.0215
12.4969 | -106.0423
12.5006 | -106.0638
12.5043 | -106.0835
12.508 | -106.0984
12.5117 | -106.1122
12.5154 | -106.1267
12.5191 | -106.1386
12.5228 | -106.1513
12.5265 | -106.1638
12.5302 | -106.1703
12.5339 | -106.1702
12.5376 | -106.1717
12.5413 | -106.1695
12.545 | -106.1604
12.5487 | -106.15
12.5524 | -106.1385
12.5561 | -106.1264
12.5598 | -106.1125
12.5635 | -106.098
12.5672 | -106.084
12.571 | -106.0718
12.5747 | -106.0637
12.5784 | -106.0532
12.5821 | -106.0432
12.5858 | -106.0319
12.5895 | -106.0223
12.5932 | -106.0149
12.5969 | -106.0109
12.6006 | -106.0093
12.6043 | -106.0077
12.608 | -106.0055
12.6117 | -106.0052
12.6154 | -106.0085
12.6191 | -106.0126
12.6228 | -106.0188
12.6265 | -106.0247
12.6302 | -106.0285
12.6339 | -106.0306
12.6376 | -106.035
12.6413 | -106.0401
12.645 | -106.0467
12.6487 | -106.0529
12.6524 | -106.0544
12.6561 | -106.0515
12.6598 | -106.051
12.6635 | -106.0474
12.6673 | -106.0388
12.671 | -106.0272
12.6747 | -106.0137
12.6784 | -105.9969
12.6821 | -105.9765
12.6858 | -105.9568
12.6895 | -105.9333
12.6932 | -105.9044
12.6969 | -105.875
12.7006 | -105.8445
12.7043 | -105.8119
12.708 | -105.7807
12.7117 | -105.7549
12.7154 | -105.7298
12.7191 | -105.706
12.7228 | -105.6884
12.7265 | -105.6746
12.7302 | -105.6615
12.7339 | -105.6535
12.7376 | -105.6508
12.7413 | -105.6536
12.745 | -105.6607
12.7487 | -105.6705
12.7524 | -105.6831
12.7561 | -105.6991
12.7599 | -105.7243
12.7636 | -105.7509
12.7673 | -105.7793
12.771 | -105.8112
12.7747 | -105.8442
12.7784 | -105.8755
12.7821 | -105.908
12.7858 | -105.9398
12.7895 | -105.9725
12.7932 | -106.002
12.7969 | -106.0294
12.8006 | -106.0549
12.8043 | -106.0775
12.808 | -106.097
12.8117 | -106.1156
12.8154 | -106.1299
12.8191 | -106.1377
12.8228 | -106.1325
12.8265 | -106.1373
12.8302 | -106.1415
12.8339 | -106.1438
12.8376 | -106.1463
12.8413 | -106.1479
12.845 | -106.1466
12.8487 | -106.1442
12.8524 | -106.1428
12.8562 | -106.1403
12.8599 | -106.1357
12.8636 | -106.1344
12.8673 | -106.1369
12.871 | -106.1392
12.8747 | -106.1431
12.8784 | -106.1497
12.8821 | -106.1576
12.8858 | -106.1686
12.8895 | -106.1818
12.8932 | -106.1953
12.8969 | -106.2116
12.9006 | -106.233
12.9043 | -106.2541
12.908 | -106.2737
12.9117 | -106.2945
12.9154 | -106.3157
12.9191 | -106.3367
12.9228 | -106.3599
12.9265 | -106.3818
12.9302 | -106.4009
12.9339 | -106.4196
12.9376 | -106.4388
12.9413 | -106.4542
12.945 | -106.4662
12.9487 | -106.48
12.9525 | -106.4923
12.9562 | -106.496
12.9599 | -106.497
12.9636 | -106.4974
12.9673 | -106.4964
12.971 | -106.4921
12.9747 | -106.482
12.9784 | -106.4719
12.9821 | -106.4602
12.9858 | -106.451
12.9895 | -106.4401
12.9932 | -106.4273
12.9969 | -106.416
13.0006 | -106.4069
13.0043 | -106.3938
13.008 | -106.3763
13.0117 | -106.3631
13.0154 | -106.3509
13.0191 | -106.3409
13.0228 | -106.332
13.0265 | -106.3225
13.0302 | -106.3148
13.0339 | -106.3089
13.0376 | -106.3017
13.0413 | -106.2942
13.045 | -106.292
13.0488 | -106.2933
13.0525 | -106.2918
13.0562 | -106.2906
13.0599 | -106.2921
13.0636 | -106.296
13.0673 | -106.2992
13.071 | -106.3003
13.0747 | -106.3042
13.0784 | -106.3057
13.0821 | -106.308
13.0858 | -106.3109
13.0895 | -106.3147
13.0932 | -106.3186
13.0969 | -106.3228
13.1006 | -106.3299
13.1043 | -106.3374
13.108 | -106.3422
13.1117 | -106.3474
13.1154 | -106.3517
13.1191 | -106.3569
13.1228 | -106.3624
13.1265 | -106.3629
13.1302 | -106.3577
13.1339 | -106.3509
13.1376 | -106.3427
13.1413 | -106.3332
13.1451 | -106.3235
13.1488 | -106.3102
13.1525 | -106.2944
13.1562 | -106.2754
13.1599 | -106.257
13.1636 | -106.239
13.1673 | -106.2154
13.171 | -106.1898
13.1747 | -106.1624
13.1784 | -106.1359
13.1821 | -106.1128
13.1858 | -106.0869
13.1895 | -106.061
13.1932 | -106.0425
13.1969 | -106.0288
13.2006 | -106.0123
13.2043 | -105.9965
13.208 | -105.9847
13.2117 | -105.9795
13.2154 | -105.977
13.2191 | -105.9768
13.2228 | -105.9842
13.2265 | -105.9958
13.2302 | -106.0099
13.2339 | -106.0216
13.2377 | -106.0355
13.2414 | -106.0517
13.2451 | -106.0689
13.2488 | -106.0867
13.2525 | -106.1046
13.2562 | -106.1263
13.2599 | -106.1474
13.2636 | -106.1666
13.2673 | -106.1864
13.271 | -106.2051
13.2747 | -106.2198
13.2784 | -106.2306
13.2821 | -106.2392
13.2858 | -106.2453
13.2895 | -106.2448
13.2932 | -106.242
13.2969 | -106.2421
13.3006 | -106.2351
13.3043 | -106.2246
13.308 | -106.2121
13.3117 | -106.1989
13.3154 | -106.1838
13.3191 | -106.1708
13.3228 | -106.1555
13.3265 | -106.1395
13.3302 | -106.1249
13.334 | -106.1071
13.3377 | -106.0919
13.3414 | -106.0795
13.3451 | -106.0657
13.3488 | -106.0551
13.3525 | -106.0453
13.3562 | -106.0465
13.3599 | -106.0405
13.3636 | -106.0366
13.3673 | -106.032
13.371 | -106.0293
13.3747 | -106.0283
13.3784 | -106.0291
13.3821 | -106.029
13.3858 | -106.0285
13.3895 | -106.0297
13.3932 | -106.033
13.3969 | -106.0378
13.4006 | -106.0386
13.4043 | -106.0348
13.408 | -106.0362
13.4117 | -106.037
13.4154 | -106.0339
13.4191 | -106.0337
13.4228 | -106.0315
13.4265 | -106.0301
13.4303 | -106.0273
13.434 | -106.0223
13.4377 | -106.019
13.4414 | -106.0171
13.4451 | -106.011
13.4488 | -106.0027
13.4525 | -105.9957
13.4562 | -105.9877
13.4599 | -105.9779
13.4636 | -105.9677
13.4673 | -105.9585
13.471 | -105.9515
13.4747 | -105.9466
13.4784 | -105.938
13.4821 | -105.9319
13.4858 | -105.9246
13.4895 | -105.9123
13.4932 | -105.9049
13.4969 | -105.8995
13.5006 | -105.8913
13.5043 | -105.8811
13.508 | -105.8747
13.5117 | -105.8679
13.5154 | -105.8632
13.5191 | -105.8551
13.5228 | -105.844
13.5266 | -105.8379
13.5303 | -105.8324
13.534 | -105.827
13.5377 | -105.8262
13.5414 | -105.8244
13.5451 | -105.8263
13.5488 | -105.8314
13.5525 | -105.8369
13.5562 | -105.843
13.5599 | -105.8507
13.5636 | -105.8597
13.5673 | -105.8681
13.571 | -105.8793
13.5747 | -105.8933
13.5784 | -105.9068
13.5821 | -105.9185
13.5858 | -105.9297
13.5895 | -105.9402
13.5932 | -105.953
13.5969 | -105.9669
13.6006 | -105.9796
13.6043 | -105.9888
13.608 | -105.9969
13.6117 | -106.0037
13.6154 | -106.0089
13.6191 | -106.0143
13.6229 | -106.0177
13.6266 | -106.0198
13.6303 | -106.021
13.634 | -106.0176
13.6377 | -106.0111
13.6414 | -106.0094
13.6451 | -106.0071
13.6488 | -106.0047
13.6525 | -106.0041
13.6562 | -106.0052
13.6599 | -106.005
13.6636 | -106.0042
13.6673 | -106.0055
13.671 | -106.0052
13.6747 | -106.0064
13.6784 | -106.0081
13.6821 | -106.0126
13.6858 | -106.017
13.6895 | -106.0184
13.6932 | -106.0241
13.6969 | -106.0283
13.7006 | -106.0304
13.7043 | -106.0311
13.708 | -106.0347
13.7117 | -106.036
13.7155 | -106.039
13.7192 | -106.0425
13.7229 | -106.0435
13.7266 | -106.043
13.7303 | -106.0374
13.734 | -106.0309
13.7377 | -106.0278
13.7414 | -106.0232
13.7451 | -106.0175
13.7488 | -106.0132
13.7525 | -106.0071
13.7562 | -106.0008
13.7599 | -105.9956
13.7636 | -105.9928
13.7673 | -105.989
13.771 | -105.9877
13.7747 | -105.9887
13.7784 | -105.9889
13.7821 | -105.9866
13.7858 | -105.9897
13.7895 | -105.9948
13.7932 | -106.0025
13.7969 | -106.0126
13.8006 | -106.0225
13.8043 | -106.0365
13.808 | -106.0508
13.8118 | -106.0626
13.8155 | -106.0777
13.8192 | -106.0943
13.8229 | -106.1111
13.8266 | -106.1302
13.8303 | -106.1487
13.834 | -106.1669
13.8377 | -106.1845
13.8414 | -106.1994
13.8451 | -106.2129
13.8488 | -106.2241
13.8525 | -106.2337
13.8562 | -106.2443
13.8599 | -106.2527
13.8636 | -106.2582
13.8673 | -106.2651
13.871 | -106.2716
13.8747 | -106.2749
13.8784 | -106.2771
13.8821 | -106.2794
13.8858 | -106.2823
13.8895 | -106.2766
13.8932 | -106.2734
13.8969 | -106.2736
13.9006 | -106.2738
13.9043 | -106.2724
13.9081 | -106.2681
13.9118 | -106.2652
13.9155 | -106.2635
13.9192 | -106.2634
13.9229 | -106.2641
13.9266 | -106.264
13.9303 | -106.2688
13.934 | -106.2728
13.9377 | -106.2762
13.9414 | -106.2814
13.9451 | -106.2891
13.9488 | -106.2981
13.9525 | -106.3123
13.9562 | -106.3247
13.9599 | -106.3248
13.9636 | -106.3219
13.9673 | -106.3271
13.971 | -106.3331
13.9747 | -106.3386
13.9784 | -106.3426
13.9821 | -106.3476
13.9858 | -106.3472
13.9895 | -106.343
13.9932 | -106.3366
13.9969 | -106.3285
14.0006 | -106.3192
14.0044 | -106.3078
14.0081 | -106.2941
14.0118 | -106.2794
14.0155 | -106.2576
14.0192 | -106.2345
14.0229 | -106.2103
14.0266 | -106.1812
14.0303 | -106.151
14.034 | -106.121
14.0377 | -106.0892
14.0414 | -106.0569
14.0451 | -106.0291
14.0488 | -106.0052
14.0525 | -105.983
14.0562 | -105.9604
14.0599 | -105.937
14.0636 | -105.9165
14.0673 | -105.8985
14.071 | -105.8849
14.0747 | -105.8732
14.0784 | -105.8646
14.0821 | -105.86
14.0858 | -105.8552
14.0895 | -105.8548
14.0932 | -105.8591
14.0969 | -105.8651
14.1007 | -105.8691
14.1044 | -105.8745
14.1081 | -105.8834
14.1118 | -105.8944
14.1155 | -105.9064
14.1192 | -105.9207
14.1229 | -105.9373
14.1266 | -105.9551
14.1303 | -105.975
14.134 | -105.9959
14.1377 | -106.0181
14.1414 | -106.0407
14.1451 | -106.0656
14.1488 | -106.0911
14.1525 | -106.1182
14.1562 | -106.146
14.1599 | -106.1764
14.1636 | -106.2171
14.1673 | -106.2536
14.171 | -106.2791
14.1747 | -106.3051
14.1784 | -106.3334
14.1821 | -106.3631
14.1858 | -106.39
14.1895 | -106.4159
14.1933 | -106.4408
14.197 | -106.4667
14.2007 | -106.4913
14.2044 | -106.5163
14.2081 | -106.5392
14.2118 | -106.5618
14.2155 | -106.5838
14.2192 | -106.6047
14.2229 | -106.624
14.2266 | -106.6423
14.2303 | -106.6605
14.234 | -106.6786
14.2377 | -106.6925
14.2414 | -106.7077
14.2451 | -106.7248
14.2488 | -106.735
14.2525 | -106.7436
14.2562 | -106.7525
14.2599 | -106.7642
14.2636 | -106.779
14.2673 | -106.7893
14.271 | -106.7985
14.2747 | -106.8068
14.2784 | -106.8121
14.2821 | -106.8175
14.2858 | -106.8248
14.2896 | -106.8342
14.2933 | -106.8398
14.297 | -106.8391
14.3007 | -106.8388
14.3044 | -106.8457
14.3081 | -106.8531
14.3118 | -106.8567
14.3155 | -106.8565
14.3192 | -106.8537
14.3229 | -106.847
14.3266 | -106.8386
14.3303 | -106.8329
14.334 | -106.8276
14.3377 | -106.8216
14.3414 | -106.8125
14.3451 | -106.7985
14.3488 | -106.7812
14.3525 | -106.7657
14.3562 | -106.7479
14.3599 | -106.7296
14.3636 | -106.7112
14.3673 | -106.6916
14.371 | -106.6737
14.3747 | -106.6563
14.3784 | -106.6369
14.3821 | -106.6168
14.3859 | -106.598
14.3896 | -106.5795
14.3933 | -106.5598
14.397 | -106.5383
14.4007 | -106.5152
14.4044 | -106.4939
14.4081 | -106.4719
14.4118 | -106.4512
14.4155 | -106.4347
14.4192 | -106.4188
14.4229 | -106.4026
14.4266 | -106.3883
14.4303 | -106.3776
14.434 | -106.366
14.4377 | -106.3538
14.4414 | -106.3433
14.4451 | -106.336
14.4488 | -106.3287
14.4525 | -106.3208
14.4562 | -106.3142
14.4599 | -106.3088
14.4636 | -106.3025
14.4673 | -106.2975
14.471 | -106.2925
14.4747 | -106.2915
14.4784 | -106.2908
14.4822 | -106.2905
14.4859 | -106.2915
14.4896 | -106.2924
14.4933 | -106.2908
14.497 | -106.2903
14.5007 | -106.2927
14.5044 | -106.2965
14.5081 | -106.2972
14.5118 | -106.2957
14.5155 | -106.2953
14.5192 | -106.2952
14.5229 | -106.2956
14.5266 | -106.2998
14.5303 | -106.3068
14.534 | -106.3137
14.5377 | -106.3172
14.5414 | -106.3204
14.5451 | -106.3237
14.5488 | -106.3284
14.5525 | -106.3367
14.5562 | -106.3468
14.5599 | -106.3541
14.5636 | -106.3598
14.5673 | -106.3637
14.571 | -106.362
14.5747 | -106.3623
14.5785 | -106.3658
14.5822 | -106.3673
14.5859 | -106.3706
14.5896 | -106.3751
14.5933 | -106.3769
14.597 | -106.3798
14.6007 | -106.3832
14.6044 | -106.3864
14.6081 | -106.3891
14.6118 | -106.3925
14.6155 | -106.3958
14.6192 | -106.398
14.6229 | -106.3985
14.6266 | -106.4004
14.6303 | -106.4053
14.634 | -106.4094
14.6377 | -106.4109
14.6414 | -106.4074
14.6451 | -106.4052
14.6488 | -106.4082
14.6525 | -106.411
14.6562 | -106.4103
14.6599 | -106.4081
14.6636 | -106.4067
14.6673 | -106.4069
14.671 | -106.4067
14.6748 | -106.4109
14.6785 | -106.4127
14.6822 | -106.4131
14.6859 | -106.4086
14.6896 | -106.4058
14.6933 | -106.4037
14.697 | -106.4021
14.7007 | -106.3997
14.7044 | -106.3977
14.7081 | -106.3952
14.7118 | -106.3927
14.7155 | -106.3893
14.7192 | -106.3884
14.7229 | -106.3845
14.7266 | -106.3774
14.7303 | -106.3721
14.734 | -106.3683
14.7377 | -106.363
14.7414 | -106.3546
14.7451 | -106.3444
14.7488 | -106.3383
14.7525 | -106.3349
14.7562 | -106.3289
14.7599 | -106.3239
14.7636 | -106.3182
14.7674 | -106.3122
14.7711 | -106.3066
14.7748 | -106.3027
14.7785 | -106.3008
14.7822 | -106.3036
14.7859 | -106.3047
14.7896 | -106.3023
14.7933 | -106.3001
14.797 | -106.2963
14.8007 | -106.2951
14.8044 | -106.2985
14.8081 | -106.3019
14.8118 | -106.3069
14.8155 | -106.3139
14.8192 | -106.3202
14.8229 | -106.3266
14.8266 | -106.3333
14.8303 | -106.3401
14.834 | -106.3459
14.8377 | -106.3534
14.8414 | -106.3642
14.8451 | -106.3751
14.8488 | -106.3852
14.8525 | -106.3957
14.8562 | -106.4059
14.8599 | -106.4153
14.8637 | -106.4234
14.8674 | -106.4322
14.8711 | -106.4428
14.8748 | -106.4548
14.8785 | -106.4657
14.8822 | -106.4767
14.8859 | -106.4848
14.8896 | -106.4924
14.8933 | -106.5019
14.897 | -106.5135
14.9007 | -106.5238
14.9044 | -106.5327
14.9081 | -106.5384
14.9118 | -106.541
14.9155 | -106.5462
14.9192 | -106.5487
14.9229 | -106.5487
14.9266 | -106.5526
14.9303 | -106.5539
14.934 | -106.5521
14.9377 | -106.5485
14.9414 | -106.5437
14.9451 | -106.5385
14.9488 | -106.5311
14.9525 | -106.5248
14.9562 | -106.5186
14.96 | -106.5091
14.9637 | -106.5009
14.9674 | -106.4891
14.9711 | -106.4803
14.9748 | -106.4716
14.9785 | -106.4624
14.9822 | -106.4531
14.9859 | -106.444
14.9896 | -106.4364
14.9933 | -106.4306
14.997 | -106.4243
15.0007 | -106.4186
15.0044 | -106.4142
15.0081 | -106.4096
15.0118 | -106.4069
15.0155 | -106.4035
15.0192 | -106.3968
15.0229 | -106.3919
15.0266 | -106.3884
15.0303 | -106.3855
15.034 | -106.3819
15.0377 | -106.3753
15.0414 | -106.3685
15.0451 | -106.3632
15.0488 | -106.3598
15.0525 | -106.3541
15.0563 | -106.3469
15.06 | -106.341
15.0637 | -106.3357
15.0674 | -106.3317
15.0711 | -106.3263
15.0748 | -106.3194
15.0785 | -106.314
15.0822 | -106.3087
15.0859 | -106.3031
15.0896 | -106.2968
15.0933 | -106.2922
15.097 | -106.2858
15.1007 | -106.2808
15.1044 | -106.278
15.1081 | -106.2768
15.1118 | -106.2744
15.1155 | -106.2695
15.1192 | -106.2667
15.1229 | -106.2654
15.1266 | -106.2639
15.1303 | -106.2622
15.134 | -106.2618
15.1377 | -106.2591
15.1414 | -106.2544
15.1451 | -106.2499
15.1488 | -106.2479
15.1526 | -106.2493
15.1563 | -106.2489
15.16 | -106.2464
15.1637 | -106.2437
15.1674 | -106.242
15.1711 | -106.2416
15.1748 | -106.2427
15.1785 | -106.2416
15.1822 | -106.2379
15.1859 | -106.2363
15.1896 | -106.2352
15.1933 | -106.2369
15.197 | -106.2379
15.2007 | -106.2362
15.2044 | -106.2337
15.2081 | -106.2327
15.2118 | -106.2303
15.2155 | -106.2275
15.2192 | -106.2255
15.2229 | -106.224
15.2266 | -106.225
15.2303 | -106.2251
15.234 | -106.2259
15.2377 | -106.2294
15.2414 | -106.2303
15.2452 | -106.231
15.2489 | -106.2325
15.2526 | -106.2371
15.2563 | -106.2437
15.26 | -106.2495
15.2637 | -106.2535
15.2674 | -106.2556
15.2711 | -106.2604
15.2748 | -106.2649
15.2785 | -106.2683
15.2822 | -106.2721
15.2859 | -106.2784
15.2896 | -106.284
15.2933 | -106.2863
15.297 | -106.2893
15.3007 | -106.2944
15.3044 | -106.2966
15.3081 | -106.2962
15.3118 | -106.2941
15.3155 | -106.2933
15.3192 | -106.2916
15.3229 | -106.2918
15.3266 | -106.2873
15.3303 | -106.2818
15.334 | -106.2772
15.3377 | -106.2714
15.3415 | -106.2677
15.3452 | -106.2641
15.3489 | -106.2575
15.3526 | -106.251
15.3563 | -106.2454
15.36 | -106.2401
15.3637 | -106.2359
15.3674 | -106.2304
15.3711 | -106.225
15.3748 | -106.2221
15.3785 | -106.2207
15.3822 | -106.2194
15.3859 | -106.2199
15.3896 | -106.2222
15.3933 | -106.2233
15.397 | -106.228
15.4007 | -106.237
15.4044 | -106.2432
15.4081 | -106.2497
15.4118 | -106.2531
15.4155 | -106.26
15.4192 | -106.2696
15.4229 | -106.2788
15.4266 | -106.285
15.4303 | -106.291
15.434 | -106.2999
15.4378 | -106.3074
15.4415 | -106.314
15.4452 | -106.3208
15.4489 | -106.3256
15.4526 | -106.3287
15.4563 | -106.3346
15.46 | -106.3418
15.4637 | -106.3481
15.4674 | -106.3535
15.4711 | -106.3557
15.4748 | -106.3607
15.4785 | -106.3645
15.4822 | -106.369
15.4859 | -106.3761
15.4896 | -106.3807
15.4933 | -106.3804
15.497 | -106.3837
15.5007 | -106.3869
15.5044 | -106.3884
15.5081 | -106.3852
15.5118 | -106.3839
15.5155 | -106.3861
15.5192 | -106.3855
15.5229 | -106.3834
15.5266 | -106.382
15.5303 | -106.3821
15.5341 | -106.3803
15.5378 | -106.3792
15.5415 | -106.3763
15.5452 | -106.3695
15.5489 | -106.3674
15.5526 | -106.3655
15.5563 | -106.362
15.56 | -106.3583
15.5637 | -106.3534
15.5674 | -106.3467
15.5711 | -106.3401
15.5748 | -106.3358
15.5785 | -106.333
15.5822 | -106.3318
15.5859 | -106.3295
15.5896 | -106.3267
15.5933 | -106.3231
15.597 | -106.3189
15.6007 | -106.3147
15.6044 | -106.313
15.6081 | -106.3091
15.6118 | -106.3051
15.6155 | -106.3036
15.6192 | -106.3026
15.6229 | -106.3007
15.6266 | -106.3004
15.6304 | -106.2981
15.6341 | -106.2981
15.6378 | -106.2973
15.6415 | -106.2957
15.6452 | -106.2943
15.6489 | -106.2939
15.6526 | -106.2932
15.6563 | -106.2905
15.66 | -106.2863
15.6637 | -106.2811
15.6674 | -106.2765
15.6711 | -106.2744
15.6748 | -106.2745
15.6785 | -106.2707
15.6822 | -106.2675
15.6859 | -106.2626
15.6896 | -106.2565
15.6933 | -106.2544
15.697 | -106.2503
15.7007 | -106.2509
15.7044 | -106.2504
15.7081 | -106.249
15.7118 | -106.2494
15.7155 | -106.2536
15.7192 | -106.2578
15.723 | -106.2604
15.7267 | -106.2623
15.7304 | -106.2641
15.7341 | -106.2696
15.7378 | -106.275
15.7415 | -106.2795
15.7452 | -106.2832
15.7489 | -106.2857
15.7526 | -106.2895
15.7563 | -106.2922
15.76 | -106.2951
15.7637 | -106.2981
15.7674 | -106.2993
15.7711 | -106.3004
15.7748 | -106.3007
15.7785 | -106.301
15.7822 | -106.3005
15.7859 | -106.2968
15.7896 | -106.2926
15.7933 | -106.2868
15.797 | -106.2812
15.8007 | -106.2707
15.8044 | -106.2621
15.8081 | -106.25
15.8118 | -106.2375
15.8155 | -106.229
15.8193 | -106.2211
15.823 | -106.2144
15.8267 | -106.2078
15.8304 | -106.1994
15.8341 | -106.1924
15.8378 | -106.1849
15.8415 | -106.1746
15.8452 | -106.162
15.8489 | -106.1502
15.8526 | -106.1418
15.8563 | -106.137
15.86 | -106.131
15.8637 | -106.1247
15.8674 | -106.1186
15.8711 | -106.1124
15.8748 | -106.106
15.8785 | -106.0973
15.8822 | -106.0889
15.8859 | -106.0829
15.8896 | -106.0781
15.8933 | -106.0702
15.897 | -106.061
15.9007 | -106.0485
15.9044 | -106.0368
15.9081 | -106.0261
15.9118 | -106.0156
15.9156 | -106.0025
15.9193 | -105.9889
15.923 | -105.9751
15.9267 | -105.9566
15.9304 | -105.9405
15.9341 | -105.9236
15.9378 | -105.9063
15.9415 | -105.8877
15.9452 | -105.8686
15.9489 | -105.8529
15.9526 | -105.8387
15.9563 | -105.8256
15.96 | -105.8115
15.9637 | -105.795
15.9674 | -105.7775
15.9711 | -105.7634
15.9748 | -105.7482
15.9785 | -105.7315
15.9822 | -105.7138
15.9859 | -105.6993
15.9896 | -105.6821
15.9933 | -105.6664
15.997 | -105.65
16.0007 | -105.6333
16.0044 | -105.6166
16.0081 | -105.5997
16.0119 | -105.5832
16.0156 | -105.5684
16.0193 | -105.5536
16.023 | -105.5366
16.0267 | -105.5219
16.0304 | -105.5059
16.0341 | -105.4909
16.0378 | -105.4748
16.0415 | -105.4603
16.0452 | -105.4464
16.0489 | -105.4303
16.0526 | -105.415
16.0563 | -105.4025
16.06 | -105.3918
16.0637 | -105.3826
16.0674 | -105.3752
16.0711 | -105.3695
16.0748 | -105.3653
16.0785 | -105.3606
16.0822 | -105.353
16.0859 | -105.3451
16.0896 | -105.3378
16.0933 | -105.3331
16.097 | -105.3307
16.1007 | -105.3291
16.1044 | -105.3276
16.1082 | -105.3278
16.1119 | -105.3235
16.1156 | -105.3202
16.1193 | -105.3149
16.123 | -105.3113
16.1267 | -105.3074
16.1304 | -105.3054
16.1341 | -105.3031
16.1378 | -105.2992
16.1415 | -105.2967
16.1452 | -105.2973
16.1489 | -105.295
16.1526 | -105.2932
16.1563 | -105.2939
16.16 | -105.2887
16.1637 | -105.285
16.1674 | -105.2856
16.1711 | -105.2856
16.1748 | -105.2862
16.1785 | -105.2857
16.1822 | -105.2862
16.1859 | -105.2883
16.1896 | -105.2905
16.1933 | -105.2955
16.197 | -105.3014
16.2007 | -105.3053
16.2045 | -105.3116
16.2082 | -105.3175
16.2119 | -105.3282
16.2156 | -105.336
16.2193 | -105.3445
16.223 | -105.3528
16.2267 | -105.361
16.2304 | -105.368
16.2341 | -105.3744
16.2378 | -105.3801
16.2415 | -105.3876
16.2452 | -105.3949
16.2489 | -105.4018
16.2526 | -105.4093
16.2563 | -105.415
16.26 | -105.4195
16.2637 | -105.4209
16.2674 | -105.4226
16.2711 | -105.423
16.2748 | -105.4211
16.2785 | -105.4184
16.2822 | -105.4157
16.2859 | -105.4137
16.2896 | -105.4103
16.2933 | -105.4061
16.2971 | -105.4018
16.3008 | -105.3988
16.3045 | -105.3977
16.3082 | -105.3952
16.3119 | -105.393
16.3156 | -105.3938
16.3193 | -105.3945
16.323 | -105.3968
16.3267 | -105.3991
16.3304 | -105.4003
16.3341 | -105.403
16.3378 | -105.4094
16.3415 | -105.418
16.3452 | -105.4299
16.3489 | -105.4424
16.3526 | -105.4578
16.3563 | -105.4737
16.36 | -105.4894
16.3637 | -105.5056
16.3674 | -105.5244
16.3711 | -105.5463
16.3748 | -105.5698
16.3785 | -105.5927
16.3822 | -105.6179
16.3859 | -105.6432
16.3896 | -105.6703
16.3934 | -105.7021
16.3971 | -105.7326
16.4008 | -105.7598
16.4045 | -105.7902
16.4082 | -105.8195
16.4119 | -105.8463
16.4156 | -105.8739
16.4193 | -105.9012
16.423 | -105.9271
16.4267 | -105.953
16.4304 | -105.9795
16.4341 | -106.0065
16.4378 | -106.0311
16.4415 | -106.0572
16.4452 | -106.0774
16.4489 | -106.0971
16.4526 | -106.1168
16.4563 | -106.1338
16.46 | -106.1503
16.4637 | -106.1643
16.4674 | -106.1776
16.4711 | -106.191
16.4748 | -106.1986
16.4785 | -106.2029
16.4822 | -106.2076
16.4859 | -106.2109
16.4897 | -106.2161
16.4934 | -106.2206
16.4971 | -106.2235
16.5008 | -106.2251
16.5045 | -106.2275
16.5082 | -106.2304
16.5119 | -106.2319
16.5156 | -106.2325
16.5193 | -106.2292
16.523 | -106.2266
16.5267 | -106.2228
16.5304 | -106.218
16.5341 | -106.2124
16.5378 | -106.2086
16.5415 | -106.2088
16.5452 | -106.2113
16.5489 | -106.2126
16.5526 | -106.2121
16.5563 | -106.2102
16.56 | -106.2096
16.5637 | -106.2079
16.5674 | -106.2076
16.5711 | -106.2078
16.5748 | -106.2048
16.5785 | -106.2017
16.5822 | -106.2025
16.586 | -106.1993
16.5897 | -106.1952
16.5934 | -106.1912
16.5971 | -106.1898
16.6008 | -106.1886
16.6045 | -106.1866
16.6082 | -106.182
16.6119 | -106.1775
16.6156 | -106.1769
16.6193 | -106.1761
16.623 | -106.1716
16.6267 | -106.1679
16.6304 | -106.1667
16.6341 | -106.166
16.6378 | -106.165
16.6415 | -106.1612
16.6452 | -106.1577
16.6489 | -106.1616
16.6526 | -106.161
16.6563 | -106.1562
16.66 | -106.1512
16.6637 | -106.1457
16.6674 | -106.14
16.6711 | -106.1341
16.6748 | -106.1291
16.6786 | -106.1229
16.6823 | -106.1168
16.686 | -106.1101
16.6897 | -106.1029
16.6934 | -106.0973
16.6971 | -106.0882
16.7008 | -106.082
16.7045 | -106.0764
16.7082 | -106.0684
16.7119 | -106.0618
16.7156 | -106.0546
16.7193 | -106.0463
16.723 | -106.0369
16.7267 | -106.0235
16.7304 | -106.0129
16.7341 | -106.0059
16.7378 | -105.9963
16.7415 | -105.9892
16.7452 | -105.982
16.7489 | -105.9739
16.7526 | -105.9681
16.7563 | -105.9641
16.76 | -105.9616
16.7637 | -105.9623
16.7674 | -105.9649
16.7711 | -105.9667
16.7749 | -105.9669
16.7786 | -105.9676
16.7823 | -105.9692
16.786 | -105.9748
16.7897 | -105.9857
16.7934 | -105.9993
16.7971 | -106.0112
16.8008 | -106.0254
16.8045 | -106.0416
16.8082 | -106.0598
16.8119 | -106.0789
16.8156 | -106.0974
16.8193 | -106.1177
16.823 | -106.1377
16.8267 | -106.1612
16.8304 | -106.1837
16.8341 | -106.2047
16.8378 | -106.2244
16.8415 | -106.2455
16.8452 | -106.2655
16.8489 | -106.2851
16.8526 | -106.3028
16.8563 | -106.3213
16.86 | -106.3388
16.8637 | -106.3565
16.8674 | -106.3684
16.8712 | -106.3783
16.8749 | -106.3882
16.8786 | -106.3961
16.8823 | -106.4062
16.886 | -106.4164
16.8897 | -106.4185
16.8934 | -106.4233
16.8971 | -106.4268
16.9008 | -106.4292
16.9045 | -106.4285
16.9082 | -106.426
16.9119 | -106.4193
16.9156 | -106.412
16.9193 | -106.401
16.923 | -106.3895
16.9267 | -106.3762
16.9304 | -106.3604
16.9341 | -106.3404
16.9378 | -106.319
16.9415 | -106.2837
16.9452 | -106.0474
16.9489 | -105.7256
16.9526 | -105.5618
16.9563 | -105.4337
16.96 | -105.2973
16.9637 | -105.1495
16.9675 | -104.992
16.9712 | -104.823
16.9749 | -104.643
16.9786 | -104.4546
16.9823 | -104.2612
16.986 | -104.0629
16.9897 | -103.8611
16.9934 | -103.6576
16.9971 | -103.4523
17.0008 | -103.247
17.0045 | -103.0431
17.0082 | -102.8429
17.0119 | -102.6495
17.0156 | -102.4638
17.0193 | -102.2854
17.023 | -102.1155
17.0267 | -101.9555
17.0304 | -101.8069
17.0341 | -101.6703
17.0378 | -101.5472
17.0415 | -101.4382
17.0452 | -101.3431
17.0489 | -101.2623
17.0526 | -101.1973
17.0563 | -101.1484
17.06 | -101.1143
17.0638 | -101.0957
17.0675 | -101.0931
17.0712 | -101.1071
17.0749 | -101.137
17.0786 | -101.1823
17.0823 | -101.2419
17.086 | -101.3178
17.0897 | -101.4091
17.0934 | -101.5142
17.0971 | -101.6865
17.1008 | -101.8085
17.1045 | -101.9423
17.1082 | -102.0876
17.1119 | -102.243
17.1156 | -102.4054
17.1193 | -102.5756
17.123 | -102.7529
17.1267 | -102.934
17.1304 | -103.1188
17.1341 | -103.3058
17.1378 | -103.4936
17.1415 | -103.6812
17.1452 | -103.8668
17.1489 | -104.0502
17.1526 | -104.2282
17.1563 | -104.4008
17.1601 | -104.5653
17.1638 | -104.7233
17.1675 | -104.8739
17.1712 | -105.015
17.1749 | -105.1463
17.1786 | -105.2686
17.1823 | -105.3821
17.186 | -105.4916
17.1897 | -105.7343
17.1934 | -106.0055
17.1971 | -106.072
17.2008 | -106.0808
17.2045 | -106.0888
17.2082 | -106.0961
17.2119 | -106.1022
17.2156 | -106.1097
17.2193 | -106.1185
17.223 | -106.1298
17.2267 | -106.1374
17.2304 | -106.1442
17.2341 | -106.1524
17.2378 | -106.1593
17.2415 | -106.1659
17.2452 | -106.171
17.2489 | -106.1797
17.2527 | -106.1889
17.2564 | -106.1963
17.2601 | -106.2029
17.2638 | -106.2094
17.2675 | -106.2168
17.2712 | -106.2256
17.2749 | -106.2321
17.2786 | -106.2391
17.2823 | -106.2443
17.286 | -106.2478
17.2897 | -106.2519
17.2934 | -106.2585
17.2971 | -106.2621
17.3008 | -106.2639
17.3045 | -106.2676
17.3082 | -106.2719
17.3119 | -106.274
17.3156 | -106.2768
17.3193 | -106.2799
17.323 | -106.2802
17.3267 | -106.2815
17.3304 | -106.2783
17.3341 | -106.2746
17.3378 | -106.2718
17.3415 | -106.2692
17.3452 | -106.265
17.349 | -106.2597
17.3527 | -106.2562
17.3564 | -106.2517
17.3601 | -106.2479
17.3638 | -106.2439
17.3675 | -106.2358
17.3712 | -106.2289
17.3749 | -106.224
17.3786 | -106.2175
17.3823 | -106.2106
17.386 | -106.205
17.3897 | -106.1991
17.3934 | -106.1925
17.3971 | -106.1843
17.4008 | -106.178
17.4045 | -106.1721
17.4082 | -106.1676
17.4119 | -106.1629
17.4156 | -106.1563
17.4193 | -106.1514
17.423 | -106.1468
17.4267 | -106.1424
17.4304 | -106.1378
17.4341 | -106.1337
17.4378 | -106.1302
17.4415 | -106.1286
17.4453 | -106.1281
17.449 | -106.1306
17.4527 | -106.1321
17.4564 | -106.1346
17.4601 | -106.1405
17.4638 | -106.1455
17.4675 | -106.1489
17.4712 | -106.1517
17.4749 | -106.1576
17.4786 | -106.1632
17.4823 | -106.1685
17.486 | -106.1748
17.4897 | -106.1847
17.4934 | -106.1955
17.4971 | -106.2041
17.5008 | -106.2118
17.5045 | -106.2191
17.5082 | -106.2255
17.5119 | -106.2329
17.5156 | -106.2397
17.5193 | -106.2431
17.523 | -106.2455
17.5267 | -106.2512
17.5304 | -106.258
17.5341 | -106.2652
17.5378 | -106.2695
17.5416 | -106.2707
17.5453 | -106.2711
17.549 | -106.2697
17.5527 | -106.268
17.5564 | -106.2653
17.5601 | -106.2593
17.5638 | -106.25
17.5675 | -106.2417
17.5712 | -106.234
17.5749 | -106.2239
17.5786 | -106.2148
17.5823 | -106.2033
17.586 | -106.1905
17.5897 | -106.176
17.5934 | -106.1628
17.5971 | -106.15
17.6008 | -106.1358
17.6045 | -106.1228
17.6082 | -106.1094
17.6119 | -106.0969
17.6156 | -106.0847
17.6193 | -106.0695
17.623 | -106.0547
17.6267 | -106.0418
17.6304 | -106.0313
17.6342 | -106.0197
17.6379 | -106.0085
17.6416 | -105.9972
17.6453 | -105.984
17.649 | -105.9737
17.6527 | -105.9644
17.6564 | -105.9565
17.6601 | -105.9497
17.6638 | -105.9426
17.6675 | -105.9348
17.6712 | -105.9288
17.6749 | -105.9272
17.6786 | -105.9265
17.6823 | -105.9264
17.686 | -105.9266
17.6897 | -105.9265
17.6934 | -105.9245
17.6971 | -105.9241
17.7008 | -105.9227
17.7045 | -105.9214
17.7082 | -105.9246
17.7119 | -105.9254
17.7156 | -105.9292
17.7193 | -105.9344
17.723 | -105.9381
17.7267 | -105.9395
17.7305 | -105.943
17.7342 | -105.9479
17.7379 | -105.954
17.7416 | -105.9604
17.7453 | -105.9664
17.749 | -105.9716
17.7527 | -105.98
17.7564 | -105.9889
17.7601 | -105.9975
17.7638 | -106.005
17.7675 | -106.0119
17.7712 | -106.0203
17.7749 | -106.0292
17.7786 | -106.0354
17.7823 | -106.043
17.786 | -106.0485
17.7897 | -106.0528
17.7934 | -106.0594
17.7971 | -106.0665
17.8008 | -106.0702
17.8045 | -106.0725
17.8082 | -106.076
17.8119 | -106.0843
17.8156 | -106.0924
17.8193 | -106.0974
17.823 | -106.1006
17.8268 | -106.1028
17.8305 | -106.1041
17.8342 | -106.1054
17.8379 | -106.1072
17.8416 | -106.1068
17.8453 | -106.1035
17.849 | -106.102
17.8527 | -106.0976
17.8564 | -106.0916
17.8601 | -106.0885
17.8638 | -106.084
17.8675 | -106.0784
17.8712 | -106.0748
17.8749 | -106.0697
17.8786 | -106.0623
17.8823 | -106.0548
17.886 | -106.0471
17.8897 | -106.0389
17.8934 | -106.0262
17.8971 | -106.0126
17.9008 | -106.0006
17.9045 | -105.9852
17.9082 | -105.9707
17.9119 | -105.9555
17.9156 | -105.94
17.9193 | -105.926
17.9231 | -105.9138
17.9268 | -105.9017
17.9305 | -105.8881
17.9342 | -105.8737
17.9379 | -105.8593
17.9416 | -105.8459
17.9453 | -105.8346
17.949 | -105.8255
17.9527 | -105.8153
17.9564 | -105.8058
17.9601 | -105.7985
17.9638 | -105.7893
17.9675 | -105.7824
17.9712 | -105.7734
17.9749 | -105.7662
17.9786 | -105.7597
17.9823 | -105.7529
17.986 | -105.748
17.9897 | -105.7425
17.9934 | -105.7344
17.9971 | -105.7262
18.0008 | -105.7184
18.0045 | -105.7133
18.0082 | -105.7082
18.0119 | -105.7053
18.0156 | -105.7014
18.0194 | -105.6983
18.0231 | -105.6936
18.0268 | -105.6871
18.0305 | -105.6841
18.0342 | -105.6811
18.0379 | -105.6724
18.0416 | -105.6634
18.0453 | -105.6575
18.049 | -105.6503
18.0527 | -105.643
18.0564 | -105.6384
18.0601 | -105.6322
18.0638 | -105.6231
18.0675 | -105.6139
18.0712 | -105.6035
18.0749 | -105.5913
18.0786 | -105.5781
18.0823 | -105.5678
18.086 | -105.5562
18.0897 | -105.544
18.0934 | -105.5329
18.0971 | -105.524
18.1008 | -105.5166
18.1045 | -105.5072
18.1082 | -105.4992
18.112 | -105.4899
18.1157 | -105.4767
18.1194 | -105.4659
18.1231 | -105.4553
18.1268 | -105.4426
18.1305 | -105.432
18.1342 | -105.4237
18.1379 | -105.4153
18.1416 | -105.4081
18.1453 | -105.4023
18.149 | -105.3982
18.1527 | -105.3971
18.1564 | -105.3955
18.1601 | -105.3934
18.1638 | -105.391
18.1675 | -105.3929
18.1712 | -105.3911
18.1749 | -105.3884
18.1786 | -105.3858
18.1823 | -105.3827
18.186 | -105.3763
18.1897 | -105.3702
18.1934 | -105.3653
18.1971 | -105.3605
18.2008 | -105.3576
18.2045 | -105.3531
18.2082 | -105.3452
18.212 | -105.3394
18.2157 | -105.3313
18.2194 | -105.3203
18.2231 | -105.3076
18.2268 | -105.2928
18.2305 | -105.2746
18.2342 | -105.2539
18.2379 | -105.2321
18.2416 | -105.2114
18.2453 | -105.1904
18.249 | -105.1666
18.2527 | -105.1441
18.2564 | -105.1161
18.2601 | -105.088
18.2638 | -105.0561
18.2675 | -105.019
18.2712 | -104.9812
18.2749 | -104.9464
18.2786 | -104.9107
18.2823 | -104.8732
18.286 | -104.8364
18.2897 | -104.7956
18.2934 | -104.7501
18.2971 | -104.7071
18.3008 | -104.6651
18.3046 | -104.6222
18.3083 | -104.5792
18.312 | -104.5362
18.3157 | -104.4913
18.3194 | -104.4477
18.3231 | -104.407
18.3268 | -104.3671
18.3305 | -104.3282
18.3342 | -104.2917
18.3379 | -104.2577
18.3416 | -104.2215
18.3453 | -104.1853
18.349 | -104.1515
18.3527 | -104.1164
18.3564 | -104.084
18.3601 | -104.0534
18.3638 | -104.0254
18.3675 | -104.0002
18.3712 | -103.9746
18.3749 | -103.9509
18.3786 | -103.9307
18.3823 | -103.9127
18.386 | -103.8958
18.3897 | -103.8761
18.3934 | -103.8569
18.3971 | -103.8396
18.4009 | -103.8225
18.4046 | -103.8067
18.4083 | -103.7885
18.412 | -103.7705
18.4157 | -103.7557
18.4194 | -103.738
18.4231 | -103.7235
18.4268 | -103.71
18.4305 | -103.6945
18.4342 | -103.677
18.4379 | -103.662
18.4416 | -103.6447
18.4453 | -103.6299
18.449 | -103.6165
18.4527 | -103.6026
18.4564 | -103.588
18.4601 | -103.5721
18.4638 | -103.554
18.4675 | -103.5379
18.4712 | -103.522
18.4749 | -103.5023
18.4786 | -103.4839
18.4823 | -103.4658
18.486 | -103.4485
18.4897 | -103.4305
18.4934 | -103.4125
18.4972 | -103.3953
18.5009 | -103.3786
18.5046 | -103.3618
18.5083 | -103.3443
18.512 | -103.327
18.5157 | -103.3081
18.5194 | -103.2915
18.5231 | -103.2755
18.5268 | -103.261
18.5305 | -103.2465
18.5342 | -103.2326
18.5379 | -103.2203
18.5416 | -103.2091
18.5453 | -103.2012
18.549 | -103.1945
18.5527 | -103.1885
18.5564 | -103.1857
18.5601 | -103.1842
18.5638 | -103.1816
18.5675 | -103.1804
18.5712 | -103.1792
18.5749 | -103.1792
18.5786 | -103.1838
18.5823 | -103.1874
18.586 | -103.188
18.5897 | -103.187
18.5935 | -103.1884
18.5972 | -103.1926
18.6009 | -103.1981
18.6046 | -103.2048
18.6083 | -103.2121
18.612 | -103.2183
18.6157 | -103.2256
18.6194 | -103.2326
18.6231 | -103.239
18.6268 | -103.2442
18.6305 | -103.2485
18.6342 | -103.256
18.6379 | -103.2625
18.6416 | -103.2669
18.6453 | -103.271
18.649 | -103.2767
18.6527 | -103.2854
18.6564 | -103.2937
18.6601 | -103.3011
18.6638 | -103.309
18.6675 | -103.3163
18.6712 | -103.3267
18.6749 | -103.3375
18.6786 | -103.3474
18.6823 | -103.3613
18.6861 | -103.3718
18.6898 | -103.3829
18.6935 | -103.3957
18.6972 | -103.4104
18.7009 | -103.4261
18.7046 | -103.4407
18.7083 | -103.4559
18.712 | -103.4697
18.7157 | -103.4837
18.7194 | -103.4976
18.7231 | -103.5134
18.7268 | -103.5309
18.7305 | -103.5477
18.7342 | -103.5626
18.7379 | -103.5814
18.7416 | -103.6014
18.7453 | -103.6179
18.749 | -103.636
18.7527 | -103.6521
18.7564 | -103.6665
18.7601 | -103.6839
18.7638 | -103.7011
18.7675 | -103.7174
18.7712 | -103.7301
18.7749 | -103.7409
18.7786 | -103.7506
18.7824 | -103.7603
18.7861 | -103.7711
18.7898 | -103.78
18.7935 | -103.7884
18.7972 | -103.7967
18.8009 | -103.8013
18.8046 | -103.8017
18.8083 | -103.8024
18.812 | -103.8036
18.8157 | -103.8005
18.8194 | -103.7965
18.8231 | -103.7946
18.8268 | -103.7936
18.8305 | -103.7929
18.8342 | -103.7898
18.8379 | -103.789
18.8416 | -103.7875
18.8453 | -103.7803
18.849 | -103.7741
18.8527 | -103.7715
18.8564 | -103.7706
18.8601 | -103.7667
18.8638 | -103.7599
18.8675 | -103.7494
18.8712 | -103.7418
18.8749 | -103.7356
18.8787 | -103.729
18.8824 | -103.7232
18.8861 | -103.7169
18.8898 | -103.7113
18.8935 | -103.7055
18.8972 | -103.702
18.9009 | -103.6966
18.9046 | -103.6896
18.9083 | -103.6811
18.912 | -103.6733
18.9157 | -103.6609
18.9194 | -103.6486
18.9231 | -103.6394
18.9268 | -103.6312
18.9305 | -103.6233
18.9342 | -103.6154
18.9379 | -103.6074
18.9416 | -103.597
18.9453 | -103.5848
18.949 | -103.5748
18.9527 | -103.5673
18.9564 | -103.5594
18.9601 | -103.5535
18.9638 | -103.5471
18.9675 | -103.5399
18.9712 | -103.5326
18.975 | -103.5224
18.9787 | -103.5133
18.9824 | -103.5041
18.9861 | -103.496
18.9898 | -103.4894
18.9935 | -103.4815
18.9972 | -103.4731
19.0009 | -103.4684
19.0046 | -103.4657
19.0083 | -103.4603
19.012 | -103.4557
19.0157 | -103.454
19.0194 | -103.4476
19.0231 | -103.4409
19.0268 | -103.435
19.0305 | -103.4289
19.0342 | -103.4199
19.0379 | -103.4086
19.0416 | -103.4026
19.0453 | -103.3949
19.049 | -103.3901
19.0527 | -103.3839
19.0564 | -103.3742
19.0601 | -103.3633
19.0638 | -103.3528
19.0675 | -103.34
19.0713 | -103.3264
19.075 | -103.3131
19.0787 | -103.2964
19.0824 | -103.2782
19.0861 | -103.2603
19.0898 | -103.2431
19.0935 | -103.2241
19.0972 | -103.2015
19.1009 | -103.1776
19.1046 | -103.1524
19.1083 | -103.1259
19.112 | -103.0991
19.1157 | -103.0732
19.1194 | -103.047
19.1231 | -103.0176
19.1268 | -102.9896
19.1305 | -102.9608
19.1342 | -102.928
19.1379 | -102.8953
19.1416 | -102.8602
19.1453 | -102.8251
19.149 | -102.7913
19.1527 | -102.7571
19.1564 | -102.7254
19.1601 | -102.6945
19.1639 | -102.663
19.1676 | -102.6312
19.1713 | -102.5994
19.175 | -102.5707
19.1787 | -102.5447
19.1824 | -102.5197
19.1861 | -102.4984
19.1898 | -102.4753
19.1935 | -102.45
19.1972 | -102.4261
19.2009 | -102.4023
19.2046 | -102.3812
19.2083 | -102.3612
19.212 | -102.3442
19.2157 | -102.3293
19.2194 | -102.3132
19.2231 | -102.2949
19.2268 | -102.2777
19.2305 | -102.264
19.2342 | -102.2581
19.2379 | -102.2493
19.2416 | -102.2413
19.2453 | -102.2355
19.249 | -102.2316
19.2527 | -102.228
19.2564 | -102.2245
19.2602 | -102.2194
19.2639 | -102.2158
19.2676 | -102.2155
19.2713 | -102.2172
19.275 | -102.2182
19.2787 | -102.2204
19.2824 | -102.224
19.2861 | -102.2272
19.2898 | -102.2329
19.2935 | -102.2429
19.2972 | -102.2551
19.3009 | -102.2678
19.3046 | -102.2791
19.3083 | -102.2903
19.312 | -102.3049
19.3157 | -102.3183
19.3194 | -102.3346
19.3231 | -102.3514
19.3268 | -102.3706
19.3305 | -102.3907
19.3342 | -102.4115
19.3379 | -102.4312
19.3416 | -102.4537
19.3453 | -102.4769
19.349 | -102.4986
19.3527 | -102.5211
19.3565 | -102.5425
19.3602 | -102.5642
19.3639 | -102.5853
19.3676 | -102.6059
19.3713 | -102.627
19.375 | -102.649
19.3787 | -102.6733
19.3824 | -102.6966
19.3861 | -102.7194
19.3898 | -102.7426
19.3935 | -102.7638
19.3972 | -102.7858
19.4009 | -102.805
19.4046 | -102.8261
19.4083 | -102.8512
19.412 | -102.8748
19.4157 | -102.8969
19.4194 | -102.9167
19.4231 | -102.9336
19.4268 | -102.949
19.4305 | -102.9617
19.4342 | -102.9732
19.4379 | -102.9866
19.4416 | -103.0001
19.4453 | -103.0125
19.449 | -103.0228
19.4528 | -103.0316
19.4565 | -103.0378
19.4602 | -103.0424
19.4639 | -103.0486
19.4676 | -103.0542
19.4713 | -103.0572
19.475 | -103.0611
19.4787 | -103.0604
19.4824 | -103.0599
19.4861 | -103.06
19.4898 | -103.0605
19.4935 | -103.061
19.4972 | -103.0622
19.5009 | -103.0618
19.5046 | -103.0595
19.5083 | -103.0611
19.512 | -103.0598
19.5157 | -103.0553
19.5194 | -103.0544
19.5231 | -103.0534
19.5268 | -103.0527
19.5305 | -103.0531
19.5342 | -103.0547
19.5379 | -103.0587
19.5416 | -103.0603
19.5453 | -103.0608
19.5491 | -103.0619
19.5528 | -103.064
19.5565 | -103.0666
19.5602 | -103.0702
19.5639 | -103.0743
19.5676 | -103.0799
19.5713 | -103.0853
19.575 | -103.0902
19.5787 | -103.0954
19.5824 | -103.1025
19.5861 | -103.1134
19.5898 | -103.1233
19.5935 | -103.1329
19.5972 | -103.1415
19.6009 | -103.1482
19.6046 | -103.1554
19.6083 | -103.1638
19.612 | -103.1735
19.6157 | -103.185
19.6194 | -103.196
19.6231 | -103.2066
19.6268 | -103.2195
19.6305 | -103.2317
19.6342 | -103.2441
19.6379 | -103.2581
19.6416 | -103.2723
19.6454 | -103.2879
19.6491 | -103.3039
19.6528 | -103.3182
19.6565 | -103.331
19.6602 | -103.3452
19.6639 | -103.3592
19.6676 | -103.3709
19.6713 | -103.3832
19.675 | -103.395
19.6787 | -103.4063
19.6824 | -103.4155
19.6861 | -103.4239
19.6898 | -103.4344
19.6935 | -103.4445
19.6972 | -103.4522
19.7009 | -103.4601
19.7046 | -103.4673
19.7083 | -103.4731
19.712 | -103.479
19.7157 | -103.4852
19.7194 | -103.4898
19.7231 | -103.4962
19.7268 | -103.5023
19.7305 | -103.5056
19.7342 | -103.51
19.738 | -103.5153
19.7417 | -103.5203
19.7454 | -103.5261
19.7491 | -103.5336
19.7528 | -103.541
19.7565 | -103.5513
19.7602 | -103.5618
19.7639 | -103.5716
19.7676 | -103.589
19.7713 | -103.5985
19.775 | -103.6106
19.7787 | -103.6233
19.7824 | -103.6377
19.7861 | -103.6496
19.7898 | -103.6607
19.7935 | -103.6722
19.7972 | -103.6872
19.8009 | -103.7011
19.8046 | -103.7159
19.8083 | -103.7316
19.812 | -103.7485
19.8157 | -103.764
19.8194 | -103.7784
19.8231 | -103.7932
19.8268 | -103.8079
19.8305 | -103.8221
19.8343 | -103.8356
19.838 | -103.8495
19.8417 | -103.8625
19.8454 | -103.8754
19.8491 | -103.8841
19.8528 | -103.8922
19.8565 | -103.9025
19.8602 | -103.9131
19.8639 | -103.9223
19.8676 | -103.9296
19.8713 | -103.9356
19.875 | -103.9411
19.8787 | -103.9477
19.8824 | -103.9555
19.8861 | -103.9625
19.8898 | -103.9705
19.8935 | -103.9793
19.8972 | -103.9877
19.9009 | -103.9954
19.9046 | -104.0039
19.9083 | -104.0114
19.912 | -104.0212
19.9157 | -104.0319
19.9194 | -104.0408
19.9231 | -104.0486
19.9268 | -104.055
19.9306 | -104.0617
19.9343 | -104.0693
19.938 | -104.0765
19.9417 | -104.0843
19.9454 | -104.0934
19.9491 | -104.1055
19.9528 | -104.1173
19.9565 | -104.1281
19.9602 | -104.1401
19.9639 | -104.1509
19.9676 | -104.1604
19.9713 | -104.1664
19.975 | -104.1725
19.9787 | -104.1823
19.9824 | -104.1926
19.9861 | -104.2004
19.9898 | -104.2078
19.9935 | -104.2157
19.9972 | -104.2197
20.0009 | -104.2234
20.0046 | -104.2254
20.0083 | -104.2296
20.012 | -104.2363
20.0157 | -104.2429
20.0194 | -104.2471
20.0231 | -104.2489
20.0269 | -104.2485
20.0306 | -104.2495
20.0343 | -104.2516
20.038 | -104.2526
20.0417 | -104.2512
20.0454 | -104.2491
20.0491 | -104.247
20.0528 | -104.2446
20.0565 | -104.2421
20.0602 | -104.239
20.0639 | -104.2346
20.0676 | -104.2299
20.0713 | -104.2269
20.075 | -104.2212
20.0787 | -104.2142
20.0824 | -104.2065
20.0861 | -104.1985
20.0898 | -104.1887
20.0935 | -104.1774
20.0972 | -104.1678
20.1009 | -104.1596
20.1046 | -104.1523
20.1083 | -104.1461
20.112 | -104.1363
20.1157 | -104.1265
20.1195 | -104.1186
20.1232 | -104.1108
20.1269 | -104.1049
20.1306 | -104.1005
20.1343 | -104.0963
20.138 | -104.0911
20.1417 | -104.085
20.1454 | -104.0815
20.1491 | -104.0777
20.1528 | -104.0756
20.1565 | -104.0766
20.1602 | -104.0776
20.1639 | -104.077
20.1676 | -104.0791
20.1713 | -104.0818
20.175 | -104.083
20.1787 | -104.0847
20.1824 | -104.0863
20.1861 | -104.0895
20.1898 | -104.0923
20.1935 | -104.0963
20.1972 | -104.0993
20.2009 | -104.0999
20.2046 | -104.1018
20.2083 | -104.1034
20.212 | -104.1063
20.2158 | -104.1089
20.2195 | -104.1115
20.2232 | -104.1127
20.2269 | -104.1139
20.2306 | -104.1145
20.2343 | -104.1156
20.238 | -104.1147
20.2417 | -104.1128
20.2454 | -104.1116
20.2491 | -104.1123
20.2528 | -104.1112
20.2565 | -104.1133
20.2602 | -104.1145
20.2639 | -104.1123
20.2676 | -104.1104
20.2713 | -104.1105
20.275 | -104.1095
20.2787 | -104.107
20.2824 | -104.106
20.2861 | -104.1036
20.2898 | -104.1039
20.2935 | -104.1029
20.2972 | -104.1003
20.3009 | -104.0968
20.3046 | -104.0944
20.3083 | -104.0926
20.3121 | -104.0906
20.3158 | -104.0884
20.3195 | -104.085
20.3232 | -104.0811
20.3269 | -104.0775
20.3306 | -104.073
20.3343 | -104.0674
20.338 | -104.063
20.3417 | -104.0604
20.3454 | -104.0576
20.3491 | -104.0535
20.3528 | -104.0525
20.3565 | -104.0501
20.3602 | -104.0457
20.3639 | -104.0423
20.3676 | -104.0386
20.3713 | -104.0316
20.375 | -104.0246
20.3787 | -104.0191
20.3824 | -104.0131
20.3861 | -104.0057
20.3898 | -103.9969
20.3935 | -103.987
20.3972 | -103.9763
20.4009 | -103.9676
20.4046 | -103.9595
20.4084 | -103.9485
20.4121 | -103.9372
20.4158 | -103.9263
20.4195 | -103.9154
20.4232 | -103.9048
20.4269 | -103.8933
20.4306 | -103.8829
20.4343 | -103.8694
20.438 | -103.8579
20.4417 | -103.8467
20.4454 | -103.8336
20.4491 | -103.8205
20.4528 | -103.8103
20.4565 | -103.7971
20.4602 | -103.7836
20.4639 | -103.7718
20.4676 | -103.7599
20.4713 | -103.7491
20.475 | -103.7398
20.4787 | -103.7289
20.4824 | -103.7159
20.4861 | -103.7059
20.4898 | -103.696
20.4935 | -103.6864
20.4972 | -103.6776
20.5009 | -103.6666
20.5047 | -103.6567
20.5084 | -103.6472
20.5121 | -103.6354
20.5158 | -103.6234
20.5195 | -103.6138
20.5232 | -103.6028
20.5269 | -103.5932
20.5306 | -103.5833
20.5343 | -103.5728
20.538 | -103.5624
20.5417 | -103.5537
20.5454 | -103.5386
20.5491 | -103.5272
20.5528 | -103.5176
20.5565 | -103.5084
20.5602 | -103.4974
20.5639 | -103.4851
20.5676 | -103.4719
20.5713 | -103.4571
20.575 | -103.4429
20.5787 | -103.4267
20.5824 | -103.4096
20.5861 | -103.391
20.5898 | -103.3717
20.5935 | -103.3529
20.5972 | -103.3315
20.601 | -103.3094
20.6047 | -103.2905
20.6084 | -103.2695
20.6121 | -103.2462
20.6158 | -103.2242
20.6195 | -103.2005
20.6232 | -103.1752
20.6269 | -103.147
20.6306 | -103.117
20.6343 | -103.0878
20.638 | -103.0575
20.6417 | -103.0297
20.6454 | -103.0016
20.6491 | -102.9726
20.6528 | -102.9435
20.6565 | -102.914
20.6602 | -102.8856
20.6639 | -102.8592
20.6676 | -102.8287
20.6713 | -102.797
20.675 | -102.7691
20.6787 | -102.74
20.6824 | -102.7099
20.6861 | -102.6815
20.6898 | -102.6537
20.6936 | -102.6288
20.6973 | -102.6043
20.701 | -102.5814
20.7047 | -102.5568
20.7084 | -102.5303
20.7121 | -102.5043
20.7158 | -102.4779
20.7195 | -102.454
20.7232 | -102.4317
20.7269 | -102.4105
20.7306 | -102.3917
20.7343 | -102.3732
20.738 | -102.355
20.7417 | -102.3377
20.7454 | -102.319
20.7491 | -102.3001
20.7528 | -102.2812
20.7565 | -102.2648
20.7602 | -102.2509
20.7639 | -102.2374
20.7676 | -102.2254
20.7713 | -102.2152
20.775 | -102.2081
20.7787 | -102.1985
20.7824 | -102.1911
20.7861 | -102.1834
20.7899 | -102.1753
20.7936 | -102.1664
20.7973 | -102.1585
20.801 | -102.1531
20.8047 | -102.1484
20.8084 | -102.1424
20.8121 | -102.1354
20.8158 | -102.1282
20.8195 | -102.1226
20.8232 | -102.1175
20.8269 | -102.112
20.8306 | -102.1065
20.8343 | -102.1016
20.838 | -102.0987
20.8417 | -102.0967
20.8454 | -102.0938
20.8491 | -102.0907
20.8528 | -102.0884
20.8565 | -102.0871
20.8602 | -102.0865
20.8639 | -102.0869
20.8676 | -102.0885
20.8713 | -102.088
20.875 | -102.0889
20.8787 | -102.0921
20.8824 | -102.0969
20.8862 | -102.1008
20.8899 | -102.1027
20.8936 | -102.1048
20.8973 | -102.1061
20.901 | -102.1052
20.9047 | -102.1082
20.9084 | -102.1108
20.9121 | -102.1124
20.9158 | -102.1145
20.9195 | -102.1193
20.9232 | -102.1277
20.9269 | -102.1331
20.9306 | -102.1348
20.9343 | -102.1378
20.938 | -102.1377
20.9417 | -102.1381
20.9454 | -102.1408
20.9491 | -102.143
20.9528 | -102.1434
20.9565 | -102.1428
20.9602 | -102.1393
20.9639 | -102.1359
20.9676 | -102.1336
20.9713 | -102.1289
20.975 | -102.1216
20.9787 | -102.115
20.9825 | -102.1032
20.9862 | -102.089
20.9899 | -102.0759
20.9936 | -102.0634
20.9973 | -102.0492
21.001 | -102.0347
21.0047 | -102.0217
21.0084 | -102.0093
21.0121 | -101.9931
21.0158 | -101.9773
21.0195 | -101.9592
21.0232 | -101.9408
21.0269 | -101.9206
21.0306 | -101.8991
21.0343 | -101.8802
21.038 | -101.8614
21.0417 | -101.8432
21.0454 | -101.8276
21.0491 | -101.8132
21.0528 | -101.7967
21.0565 | -101.7796
21.0602 | -101.7641
21.0639 | -101.7471
21.0676 | -101.7317
21.0713 | -101.7173
21.0751 | -101.7062
21.0788 | -101.6974
21.0825 | -101.6893
21.0862 | -101.6836
21.0899 | -101.6811
21.0936 | -101.6791
21.0973 | -101.6767
21.101 | -101.6769
21.1047 | -101.6796
21.1084 | -101.6839
21.1121 | -101.6888
21.1158 | -101.6951
21.1195 | -101.7037
21.1232 | -101.7145
21.1269 | -101.7273
21.1306 | -101.7421
21.1343 | -101.757
21.138 | -101.7742
21.1417 | -101.7915
21.1454 | -101.8085
21.1491 | -101.8286
21.1528 | -101.849
21.1565 | -101.8708
21.1602 | -101.8937
21.1639 | -101.9145
21.1676 | -101.9345
21.1714 | -101.9562
21.1751 | -101.9776
21.1788 | -101.9987
21.1825 | -102.0197
21.1862 | -102.0413
21.1899 | -102.0592
21.1936 | -102.0774
21.1973 | -102.0965
21.201 | -102.1141
21.2047 | -102.1289
21.2084 | -102.1419
21.2121 | -102.1531
21.2158 | -102.1636
21.2195 | -102.1707
21.2232 | -102.1733
21.2269 | -102.1767
21.2306 | -102.1784
21.2343 | -102.176
21.238 | -102.1728
21.2417 | -102.164
21.2454 | -102.1523
21.2491 | -102.1399
21.2528 | -102.128
21.2565 | -102.1133
21.2602 | -102.0953
21.2639 | -102.0794
21.2677 | -102.0586
21.2714 | -102.0327
21.2751 | -102.0077
21.2788 | -101.9808
21.2825 | -101.9578
21.2862 | -101.9336
21.2899 | -101.9026
21.2936 | -101.8691
21.2973 | -101.833
21.301 | -101.797
21.3047 | -101.7606
21.3084 | -101.7221
21.3121 | -101.6904
21.3158 | -101.6565
21.3195 | -101.6204
21.3232 | -101.5851
21.3269 | -101.5519
21.3306 | -101.5179
21.3343 | -101.485
21.338 | -101.4507
21.3417 | -101.4132
21.3454 | -101.3768
21.3491 | -101.3421
21.3528 | -101.3079
21.3565 | -101.2742
21.3602 | -101.2415
21.364 | -101.2108
21.3677 | -101.1788
21.3714 | -101.1511
21.3751 | -101.1193
21.3788 | -101.0883
21.3825 | -101.0564
21.3862 | -101.0268
21.3899 | -100.9985
21.3936 | -100.9707
21.3973 | -100.9429
21.401 | -100.915
21.4047 | -100.8875
21.4084 | -100.8623
21.4121 | -100.8378
21.4158 | -100.8141
21.4195 | -100.79
21.4232 | -100.7666
21.4269 | -100.7435
21.4306 | -100.7217
21.4343 | -100.7005
21.438 | -100.6806
21.4417 | -100.6607
21.4454 | -100.6416
21.4491 | -100.6249
21.4528 | -100.6082
21.4565 | -100.5902
21.4603 | -100.5735
21.464 | -100.5599
21.4677 | -100.5481
21.4714 | -100.5377
21.4751 | -100.5271
21.4788 | -100.516
21.4825 | -100.5048
21.4862 | -100.4951
21.4899 | -100.4839
21.4936 | -100.4736
21.4973 | -100.4609
21.501 | -100.4489
21.5047 | -100.436
21.5084 | -100.4188
21.5121 | -100.4038
21.5158 | -100.3893
21.5195 | -100.375
21.5232 | -100.3638
21.5269 | -100.3485
21.5306 | -100.3302
21.5343 | -100.3123
21.538 | -100.2962
21.5417 | -100.2802
21.5454 | -100.2631
21.5491 | -100.2448
21.5528 | -100.2259
21.5566 | -100.2063
21.5603 | -100.1875
21.564 | -100.168
21.5677 | -100.146
21.5714 | -100.1246
21.5751 | -100.1041
21.5788 | -100.0841
21.5825 | -100.0634
21.5862 | -100.0399
21.5899 | -100.0167
21.5936 | -99.995
21.5973 | -99.9747
21.601 | -99.9561
21.6047 | -99.9362
21.6084 | -99.9174
21.6121 | -99.9005
21.6158 | -99.8792
21.6195 | -99.8579
21.6232 | -99.8389
21.6269 | -99.8235
21.6306 | -99.8054
21.6343 | -99.7884
21.638 | -99.7726
21.6417 | -99.7604
21.6454 | -99.7529
21.6491 | -99.7454
21.6529 | -99.7364
21.6566 | -99.7235
21.6603 | -99.711
21.664 | -99.7035
21.6677 | -99.7002
21.6714 | -99.6984
21.6751 | -99.6966
21.6788 | -99.6925
21.6825 | -99.6928
21.6862 | -99.694
21.6899 | -99.6983
21.6936 | -99.7019
21.6973 | -99.7033
21.701 | -99.7064
21.7047 | -99.7127
21.7084 | -99.7183
21.7121 | -99.7246
21.7158 | -99.7298
21.7195 | -99.7353
21.7232 | -99.7423
21.7269 | -99.7493
21.7306 | -99.7551
21.7343 | -99.7609
21.738 | -99.7655
21.7417 | -99.7695
21.7455 | -99.7717
21.7492 | -99.7754
21.7529 | -99.7802
21.7566 | -99.7813
21.7603 | -99.7807
21.764 | -99.7819
21.7677 | -99.7828
21.7714 | -99.7825
21.7751 | -99.781
21.7788 | -99.7793
21.7825 | -99.7789
21.7862 | -99.7755
21.7899 | -99.7713
21.7936 | -99.7674
21.7973 | -99.7668
21.801 | -99.7678
21.8047 | -99.7667
21.8084 | -99.7623
21.8121 | -99.7612
21.8158 | -99.7629
21.8195 | -99.7628
21.8232 | -99.7615
21.8269 | -99.7605
21.8306 | -99.7575
21.8343 | -99.7579
21.838 | -99.7589
21.8418 | -99.7589
21.8455 | -99.7579
21.8492 | -99.7521
21.8529 | -99.7484
21.8566 | -99.7471
21.8603 | -99.7473
21.864 | -99.748
21.8677 | -99.7492
21.8714 | -99.7524
21.8751 | -99.7556
21.8788 | -99.7577
21.8825 | -99.7606
21.8862 | -99.7627
21.8899 | -99.7667
21.8936 | -99.7732
21.8973 | -99.7778
21.901 | -99.7822
21.9047 | -99.7879
21.9084 | -99.7955
21.9121 | -99.8004
21.9158 | -99.8083
21.9195 | -99.8143
21.9232 | -99.8197
21.9269 | -99.8273
21.9306 | -99.8359
21.9343 | -99.844
21.9381 | -99.85
21.9418 | -99.8595
21.9455 | -99.8698
21.9492 | -99.8782
21.9529 | -99.8876
21.9566 | -99.8945
21.9603 | -99.901
21.964 | -99.911
21.9677 | -99.9262
21.9714 | -99.9407
21.9751 | -99.9517
21.9788 | -99.9611
21.9825 | -99.972
21.9862 | -99.9841
21.9899 | -99.999
21.9936 | -100.0132
21.9973 | -100.0271
22.001 | -100.0385
22.0047 | -100.0532
22.0084 | -100.0727
22.0121 | -100.091
22.0158 | -100.1076
22.0195 | -100.1259
22.0232 | -100.1428
22.0269 | -100.16
22.0307 | -100.1785
22.0344 | -100.197
22.0381 | -100.2157
22.0418 | -100.2375
22.0455 | -100.2582
22.0492 | -100.2791
22.0529 | -100.301
22.0566 | -100.3219
22.0603 | -100.3409
22.064 | -100.3588
22.0677 | -100.3754
22.0714 | -100.3928
22.0751 | -100.4074
22.0788 | -100.4227
22.0825 | -100.4373
22.0862 | -100.4529
22.0899 | -100.4688
22.0936 | -100.4827
22.0973 | -100.4986
22.101 | -100.5127
22.1047 | -100.5248
22.1084 | -100.5336
22.1121 | -100.5424
22.1158 | -100.5556
22.1195 | -100.567
22.1232 | -100.5731
22.127 | -100.5779
22.1307 | -100.583
22.1344 | -100.5886
22.1381 | -100.5927
22.1418 | -100.5954
22.1455 | -100.5984
22.1492 | -100.6046
22.1529 | -100.6113
22.1566 | -100.6159
22.1603 | -100.6229
22.164 | -100.6303
22.1677 | -100.6348
22.1714 | -100.6349
22.1751 | -100.6349
22.1788 | -100.6348
22.1825 | -100.6317
22.1862 | -100.6282
22.1899 | -100.6289
22.1936 | -100.6307
22.1973 | -100.6347
22.201 | -100.6367
22.2047 | -100.6379
22.2084 | -100.6428
22.2121 | -100.6502
22.2158 | -100.6559
22.2195 | -100.6621
22.2233 | -100.6727
22.227 | -100.6813
22.2307 | -100.6877
22.2344 | -100.6953
22.2381 | -100.7061
22.2418 | -100.7177
22.2455 | -100.7308
22.2492 | -100.7451
22.2529 | -100.7591
22.2566 | -100.7716
22.2603 | -100.7851
22.264 | -100.7992
22.2677 | -100.8103
22.2714 | -100.8257
22.2751 | -100.8432
22.2788 | -100.8593
22.2825 | -100.8748
22.2862 | -100.8905
22.2899 | -100.9066
22.2936 | -100.9216
22.2973 | -100.9354
22.301 | -100.9506
22.3047 | -100.9666
22.3084 | -100.9814
22.3121 | -100.9967
22.3158 | -101.0126
22.3196 | -101.0272
22.3233 | -101.0405
22.327 | -101.0541
22.3307 | -101.0674
22.3344 | -101.0837
22.3381 | -101.1034
22.3418 | -101.1223
22.3455 | -101.1407
22.3492 | -101.158
22.3529 | -101.175
22.3566 | -101.1898
22.3603 | -101.207
22.364 | -101.2248
22.3677 | -101.2421
22.3714 | -101.2627
22.3751 | -101.2817
22.3788 | -101.2975
22.3825 | -101.3176
22.3862 | -101.3355
22.3899 | -101.3552
22.3936 | -101.3714
22.3973 | -101.386
22.401 | -101.3987
22.4047 | -101.4092
22.4084 | -101.4211
22.4121 | -101.4351
22.4159 | -101.4509
22.4196 | -101.4645
22.4233 | -101.4799
22.427 | -101.4929
22.4307 | -101.5052
22.4344 | -101.5196
22.4381 | -101.5352
22.4418 | -101.5488
22.4455 | -101.5635
22.4492 | -101.5761
22.4529 | -101.589
22.4566 | -101.6027
22.4603 | -101.6193
22.464 | -101.6367
22.4677 | -101.6503
22.4714 | -101.6628
22.4751 | -101.6779
22.4788 | -101.6931
22.4825 | -101.7057
22.4862 | -101.7183
22.4899 | -101.7345
22.4936 | -101.7508
22.4973 | -101.7661
22.501 | -101.7855
22.5047 | -101.8042
22.5084 | -101.8189
22.5122 | -101.8313
22.5159 | -101.8411
22.5196 | -101.8526
22.5233 | -101.8655
22.527 | -101.8761
22.5307 | -101.8841
22.5344 | -101.8933
22.5381 | -101.9013
22.5418 | -101.9088
22.5455 | -101.9162
22.5492 | -101.9249
22.5529 | -101.9334
22.5566 | -101.9372
22.5603 | -101.9401
22.564 | -101.9427
22.5677 | -101.944
22.5714 | -101.946
22.5751 | -101.9469
22.5788 | -101.9437
22.5825 | -101.943
22.5862 | -101.9446
22.5899 | -101.9428
22.5936 | -101.9403
22.5973 | -101.9346
22.601 | -101.927
22.6047 | -101.9226
22.6085 | -101.916
22.6122 | -101.9098
22.6159 | -101.9055
22.6196 | -101.8988
22.6233 | -101.8911
22.627 | -101.8817
22.6307 | -101.8719
22.6344 | -101.8623
22.6381 | -101.8508
22.6418 | -101.8368
22.6455 | -101.8231
22.6492 | -101.8112
22.6529 | -101.7967
22.6566 | -101.7801
22.6603 | -101.7661
22.664 | -101.7518
22.6677 | -101.7377
22.6714 | -101.7219
22.6751 | -101.7062
22.6788 | -101.6901
22.6825 | -101.6749
22.6862 | -101.6597
22.6899 | -101.6458
22.6936 | -101.6336
22.6973 | -101.6241
22.7011 | -101.6122
22.7048 | -101.6029
22.7085 | -101.5899
22.7122 | -101.5765
22.7159 | -101.5649
22.7196 | -101.5552
22.7233 | -101.5448
22.727 | -101.5327
22.7307 | -101.5226
22.7344 | -101.5117
22.7381 | -101.5028
22.7418 | -101.4908
22.7455 | -101.4798
22.7492 | -101.4708
22.7529 | -101.4616
22.7566 | -101.4528
22.7603 | -101.4445
22.764 | -101.4387
22.7677 | -101.4313
22.7714 | -101.4232
22.7751 | -101.4122
22.7788 | -101.3981
22.7825 | -101.3814
22.7862 | -101.3636
22.7899 | -101.3447
22.7936 | -101.3272
22.7974 | -101.3054
22.8011 | -101.2851
22.8048 | -101.2666
22.8085 | -101.247
22.8122 | -101.225
22.8159 | -101.203
22.8196 | -101.1813
22.8233 | -101.1593
22.827 | -101.1358
22.8307 | -101.1092
22.8344 | -101.0798
22.8381 | -101.0492
22.8418 | -101.0195
22.8455 | -100.9885
22.8492 | -100.9583
22.8529 | -100.9276
22.8566 | -100.8973
22.8603 | -100.8676
22.864 | -100.8375
22.8677 | -100.8056
22.8714 | -100.7732
22.8751 | -100.7423
22.8788 | -100.712
22.8825 | -100.6813
22.8862 | -100.6487
22.8899 | -100.6171
22.8937 | -100.5857
22.8974 | -100.5558
22.9011 | -100.5267
22.9048 | -100.496
22.9085 | -100.4636
22.9122 | -100.4351
22.9159 | -100.4088
22.9196 | -100.3841
22.9233 | -100.3622
22.927 | -100.339
22.9307 | -100.3167
22.9344 | -100.2963
22.9381 | -100.2775
22.9418 | -100.2603
22.9455 | -100.2439
22.9492 | -100.2293
22.9529 | -100.2149
22.9566 | -100.2025
22.9603 | -100.1925
22.964 | -100.1832
22.9677 | -100.1752
22.9714 | -100.1686
22.9751 | -100.1703
22.9788 | -100.1637
22.9825 | -100.1574
22.9862 | -100.1502
22.99 | -100.1446
22.9937 | -100.141
22.9974 | -100.1411
23.0011 | -100.1409
23.0048 | -100.1402
23.0085 | -100.1402
23.0122 | -100.1417
23.0159 | -100.1428
23.0196 | -100.1452
23.0233 | -100.1489
23.027 | -100.1516
23.0307 | -100.1506
23.0344 | -100.1527
23.0381 | -100.1572
23.0418 | -100.1618
23.0455 | -100.1665
23.0492 | -100.1694
23.0529 | -100.1734
23.0566 | -100.1739
23.0603 | -100.1773
23.064 | -100.1821
23.0677 | -100.1866
23.0714 | -100.1918
23.0751 | -100.195
23.0788 | -100.1987
23.0826 | -100.2029
23.0863 | -100.2063
23.09 | -100.2077
23.0937 | -100.2115
23.0974 | -100.216
23.1011 | -100.221
23.1048 | -100.2282
23.1085 | -100.2351
23.1122 | -100.2424
23.1159 | -100.2504
23.1196 | -100.2565
23.1233 | -100.2621
23.127 | -100.268
23.1307 | -100.2743
23.1344 | -100.2802
23.1381 | -100.2865
23.1418 | -100.2909
23.1455 | -100.2948
23.1492 | -100.2999
23.1529 | -100.3064
23.1566 | -100.314
23.1603 | -100.3211
23.164 | -100.3294
23.1677 | -100.3365
23.1714 | -100.3443
23.1751 | -100.3524
23.1789 | -100.3617
23.1826 | -100.3693
23.1863 | -100.3776
23.19 | -100.3876
23.1937 | -100.4016
23.1974 | -100.4148
23.2011 | -100.4255
23.2048 | -100.4362
23.2085 | -100.4478
23.2122 | -100.461
23.2159 | -100.4744
23.2196 | -100.4887
23.2233 | -100.5041
23.227 | -100.5211
23.2307 | -100.5387
23.2344 | -100.5562
23.2381 | -100.5701
23.2418 | -100.5827
23.2455 | -100.5976
23.2492 | -100.6127
23.2529 | -100.6281
23.2566 | -100.6444
23.2603 | -100.6601
23.264 | -100.6745
23.2677 | -100.6887
23.2714 | -100.7039
23.2752 | -100.72
23.2789 | -100.7374
23.2826 | -100.7544
23.2863 | -100.7714
23.29 | -100.7903
23.2937 | -100.8092
23.2974 | -100.826
23.3011 | -100.8431
23.3048 | -100.8626
23.3085 | -100.8808
23.3122 | -100.9007
23.3159 | -100.9211
23.3196 | -100.9406
23.3233 | -100.9578
23.327 | -100.974
23.3307 | -100.9921
23.3344 | -101.0114
23.3381 | -101.0289
23.3418 | -101.0452
23.3455 | -101.0636
23.3492 | -101.0815
23.3529 | -101.0976
23.3566 | -101.117
23.3603 | -101.1385
23.364 | -101.16
23.3677 | -101.1796
23.3715 | -101.1992
23.3752 | -101.2183
23.3789 | -101.2408
23.3826 | -101.262
23.3863 | -101.2838
23.39 | -101.3012
23.3937 | -101.3159
23.3974 | -101.3334
23.4011 | -101.3501
23.4048 | -101.3657
23.4085 | -101.3801
23.4122 | -101.3917
23.4159 | -101.405
23.4196 | -101.4181
23.4233 | -101.429
23.427 | -101.4409
23.4307 | -101.4498
23.4344 | -101.456
23.4381 | -101.4585
23.4418 | -101.4589
23.4455 | -101.462
23.4492 | -101.4667
23.4529 | -101.4716
23.4566 | -101.4742
23.4603 | -101.4737
23.464 | -101.4728
23.4678 | -101.4711
23.4715 | -101.4719
23.4752 | -101.4722
23.4789 | -101.4706
23.4826 | -101.4673
23.4863 | -101.463
23.49 | -101.4627
23.4937 | -101.459
23.4974 | -101.454
23.5011 | -101.449
23.5048 | -101.4412
23.5085 | -101.4305
23.5122 | -101.4225
23.5159 | -101.4144
23.5196 | -101.4066
23.5233 | -101.3976
23.527 | -101.3864
23.5307 | -101.3769
23.5344 | -101.368
23.5381 | -101.3588
23.5418 | -101.3511
23.5455 | -101.3411
23.5492 | -101.3308
23.5529 | -101.3212
23.5566 | -101.3114
23.5603 | -101.3005
23.5641 | -101.2903
23.5678 | -101.2823
23.5715 | -101.2757
23.5752 | -101.2701
23.5789 | -101.2639
23.5826 | -101.2578
23.5863 | -101.2523
23.59 | -101.2479
23.5937 | -101.244
23.5974 | -101.2432
23.6011 | -101.244
23.6048 | -101.2444
23.6085 | -101.2456
23.6122 | -101.2491
23.6159 | -101.2544
23.6196 | -101.2584
23.6233 | -101.2626
23.627 | -101.2687
23.6307 | -101.2772
23.6344 | -101.2862
23.6381 | -101.2967
23.6418 | -101.3099
23.6455 | -101.3218
23.6492 | -101.3346
23.6529 | -101.3492
23.6567 | -101.3664
23.6604 | -101.3816
23.6641 | -101.3958
23.6678 | -101.4119
23.6715 | -101.4286
23.6752 | -101.4453
23.6789 | -101.4614
23.6826 | -101.4782
23.6863 | -101.4943
23.69 | -101.5093
23.6937 | -101.5265
23.6974 | -101.5424
23.7011 | -101.5577
23.7048 | -101.5719
23.7085 | -101.5852
23.7122 | -101.599
23.7159 | -101.6133
23.7196 | -101.6282
23.7233 | -101.643
23.727 | -101.6584
23.7307 | -101.672
23.7344 | -101.6828
23.7381 | -101.6933
23.7418 | -101.7036
23.7455 | -101.7136
23.7492 | -101.7212
23.753 | -101.7298
23.7567 | -101.7369
23.7604 | -101.7393
23.7641 | -101.7436
23.7678 | -101.7501
23.7715 | -101.7579
23.7752 | -101.7628
23.7789 | -101.7621
23.7826 | -101.7602
23.7863 | -101.7588
23.79 | -101.7564
23.7937 | -101.752
23.7974 | -101.7466
23.8011 | -101.7428
23.8048 | -101.7381
23.8085 | -101.7328
23.8122 | -101.726
23.8159 | -101.7208
23.8196 | -101.7152
23.8233 | -101.7077
23.827 | -101.6998
23.8307 | -101.6912
23.8344 | -101.6838
23.8381 | -101.6782
23.8418 | -101.6717
23.8455 | -101.6655
23.8493 | -101.6585
23.853 | -101.6524
23.8567 | -101.6467
23.8604 | -101.6406
23.8641 | -101.6342
23.8678 | -101.6256
23.8715 | -101.62
23.8752 | -101.6143
23.8789 | -101.6072
23.8826 | -101.6009
23.8863 | -101.5937
23.89 | -101.5878
23.8937 | -101.5824
23.8974 | -101.5789
23.9011 | -101.5733
23.9048 | -101.5695
23.9085 | -101.5646
23.9122 | -101.559
23.9159 | -101.5531
23.9196 | -101.5484
23.9233 | -101.5444
23.927 | -101.5405
23.9307 | -101.5364
23.9344 | -101.5302
23.9381 | -101.5246
23.9418 | -101.5205
23.9456 | -101.5165
23.9493 | -101.512
23.953 | -101.5074
23.9567 | -101.5039
23.9604 | -101.4998
23.9641 | -101.4952
23.9678 | -101.492
23.9715 | -101.4887
23.9752 | -101.4839
23.9789 | -101.4775
23.9826 | -101.4738
23.9863 | -101.4711
23.99 | -101.4672
23.9937 | -101.4626
23.9974 | -101.4614
24.0011 | -101.4595
24.0048 | -101.4557
24.0085 | -101.4519
24.0122 | -101.4473
24.0159 | -101.4413
24.0196 | -101.4367
24.0233 | -101.4345
24.027 | -101.4333
24.0307 | -101.4318
24.0344 | -101.4295
24.0382 | -101.4271
24.0419 | -101.4265
24.0456 | -101.4285
24.0493 | -101.4292
24.053 | -101.4274
24.0567 | -101.4247
24.0604 | -101.4232
24.0641 | -101.4226
24.0678 | -101.4214
24.0715 | -101.4193
24.0752 | -101.4194
24.0789 | -101.4197
24.0826 | -101.4214
24.0863 | -101.4248
24.09 | -101.4264
24.0937 | -101.4271
24.0974 | -101.4272
24.1011 | -101.4281
24.1048 | -101.4317
24.1085 | -101.4328
24.1122 | -101.4316
24.1159 | -101.4315
24.1196 | -101.4312
24.1233 | -101.4297
24.127 | -101.4288
24.1307 | -101.4285
24.1345 | -101.4312
24.1382 | -101.4356
24.1419 | -101.438
24.1456 | -101.4405
24.1493 | -101.4437
24.153 | -101.4474
24.1567 | -101.4489
24.1604 | -101.4502
24.1641 | -101.451
24.1678 | -101.4536
24.1715 | -101.4554
24.1752 | -101.4553
24.1789 | -101.4557
24.1826 | -101.4573
24.1863 | -101.4574
24.19 | -101.4595
24.1937 | -101.4595
24.1974 | -101.4574
24.2011 | -101.4546
24.2048 | -101.4537
24.2085 | -101.4533
24.2122 | -101.4508
24.2159 | -101.4507
24.2196 | -101.4497
24.2233 | -101.448
24.227 | -101.446
24.2308 | -101.4395
24.2345 | -101.4313
24.2382 | -101.4238
24.2419 | -101.4144
24.2456 | -101.4028
24.2493 | -101.3933
24.253 | -101.3848
24.2567 | -101.3765
24.2604 | -101.3662
24.2641 | -101.3559
24.2678 | -101.3441
24.2715 | -101.3315
24.2752 | -101.3209
24.2789 | -101.3105
24.2826 | -101.3
24.2863 | -101.2873
24.29 | -101.2729
24.2937 | -101.2565
24.2974 | -101.2409
24.3011 | -101.2276
24.3048 | -101.2152
24.3085 | -101.2008
24.3122 | -101.1883
24.3159 | -101.1791
24.3196 | -101.1711
24.3233 | -101.1615
24.3271 | -101.151
24.3308 | -101.1418
24.3345 | -101.1342
24.3382 | -101.1262
24.3419 | -101.1195
24.3456 | -101.1124
24.3493 | -101.106
24.353 | -101.1027
24.3567 | -101.101
24.3604 | -101.098
24.3641 | -101.0959
24.3678 | -101.0938
24.3715 | -101.0931
24.3752 | -101.094
24.3789 | -101.0953
24.3826 | -101.097
24.3863 | -101.1001
24.39 | -101.1041
24.3937 | -101.1103
24.3974 | -101.1199
24.4011 | -101.1304
24.4048 | -101.1405
24.4085 | -101.1523
24.4122 | -101.1654
24.4159 | -101.1818
24.4196 | -101.1966
24.4234 | -101.2123
24.4271 | -101.2284
24.4308 | -101.2452
24.4345 | -101.2642
24.4382 | -101.2845
24.4419 | -101.307
24.4456 | -101.3309
24.4493 | -101.3516
24.453 | -101.3743
24.4567 | -101.4005
24.4604 | -101.4264
24.4641 | -101.4534
24.4678 | -101.4811
24.4715 | -101.51
24.4752 | -101.5386
24.4789 | -101.5641
24.4826 | -101.5897
24.4863 | -101.6176
24.49 | -101.6444
24.4937 | -101.6724
24.4974 | -101.7013
24.5011 | -101.7319
24.5048 | -101.7634
24.5085 | -101.7944
24.5122 | -101.8256
24.516 | -101.8569
24.5197 | -101.8867
24.5234 | -101.917
24.5271 | -101.945
24.5308 | -101.9729
24.5345 | -102.0019
24.5382 | -102.0316
24.5419 | -102.0624
24.5456 | -102.0935
24.5493 | -102.1228
24.553 | -102.1509
24.5567 | -102.1775
24.5604 | -102.202
24.5641 | -102.2268
24.5678 | -102.2503
24.5715 | -102.2711
24.5752 | -102.2905
24.5789 | -102.2974
24.5826 | -102.3118
24.5863 | -102.3085
24.59 | -102.2797
24.5937 | -102.2422
24.5974 | -102.2097
24.6011 | -102.1856
24.6048 | -102.1643
24.6085 | -102.1407
24.6123 | -102.1099
24.616 | -102.074
24.6197 | -102.0358
24.6234 | -101.9953
24.6271 | -101.9511
24.6308 | -101.9027
24.6345 | -101.8542
24.6382 | -101.8021
24.6419 | -101.7487
24.6456 | -101.693
24.6493 | -101.6321
24.653 | -101.565
24.6567 | -101.4965
24.6604 | -101.4264
24.6641 | -101.3515
24.6678 | -101.2744
24.6715 | -101.1973
24.6752 | -101.1199
24.6789 | -101.041
24.6826 | -100.9613
24.6863 | -100.8815
24.69 | -100.8025
24.6937 | -100.7248
24.6974 | -100.6482
24.7011 | -100.572
24.7048 | -100.4975
24.7086 | -100.4247
24.7123 | -100.3539
24.716 | -100.2868
24.7197 | -100.2219
24.7234 | -100.16
24.7271 | -100.102
24.7308 | -100.048
24.7345 | -99.9985
24.7382 | -99.953
24.7419 | -99.912
24.7456 | -99.8756
24.7493 | -99.8442
24.753 | -99.8178
24.7567 | -99.7963
24.7604 | -99.7805
24.7641 | -99.7696
24.7678 | -99.7648
24.7715 | -99.7659
24.7752 | -99.772
24.7789 | -99.7837
24.7826 | -99.8011
24.7863 | -99.8237
24.79 | -99.8519
24.7937 | -99.8857
24.7974 | -99.9257
24.8011 | -99.9706
24.8049 | -100.0206
24.8086 | -100.0756
24.8123 | -100.136
24.816 | -100.201
24.8197 | -100.2699
24.8234 | -100.3438
24.8271 | -100.4226
24.8308 | -100.5053
24.8345 | -100.5906
24.8382 | -100.6803
24.8419 | -100.7734
24.8456 | -100.8691
24.8493 | -100.9675
24.853 | -101.0677
24.8567 | -101.1704
24.8604 | -101.2727
24.8641 | -101.3761
24.8678 | -101.4815
24.8715 | -101.586
24.8752 | -101.6914
24.8789 | -101.7962
24.8826 | -101.8997
24.8863 | -102.0026
24.89 | -102.1029
24.8937 | -102.2012
24.8974 | -102.2981
24.9012 | -102.3938
24.9049 | -102.4877
24.9086 | -102.5788
24.9123 | -102.6655
24.916 | -102.7491
24.9197 | -102.8318
24.9234 | -102.9122
24.9271 | -102.9874
24.9308 | -103.0595
24.9345 | -103.1316
24.9382 | -103.2155
24.9419 | -103.321
24.9456 | -103.4215
24.9493 | -103.4983
24.953 | -103.5492
24.9567 | -103.5881
24.9604 | -103.6222
24.9641 | -103.6569
24.9678 | -103.6917
24.9715 | -103.7219
24.9752 | -103.7438
24.9789 | -103.7616
24.9826 | -103.7796
24.9863 | -103.7928
24.99 | -103.8054
24.9937 | -103.8163
24.9975 | -103.8243
25.0012 | -103.8335
25.0049 | -103.8437
25.0086 | -103.8524
25.0123 | -103.8575
25.016 | -103.8634
25.0197 | -103.8705
25.0234 | -103.8758
25.0271 | -103.8797
25.0308 | -103.8852
25.0345 | -103.8898
25.0382 | -103.8925
25.0419 | -103.8957
25.0456 | -103.9009
25.0493 | -103.9041
25.053 | -103.9049
25.0567 | -103.9049
25.0604 | -103.9054
25.0641 | -103.9061
25.0678 | -103.9045
25.0715 | -103.9033
25.0752 | -103.9024
25.0789 | -103.9023
25.0826 | -103.9007
25.0863 | -103.9
25.0901 | -103.8974
25.0938 | -103.8933
25.0975 | -103.8873
25.1012 | -103.8814
25.1049 | -103.8775
25.1086 | -103.8727
25.1123 | -103.8652
25.116 | -103.8616
25.1197 | -103.8566
25.1234 | -103.8525
25.1271 | -103.8472
25.1308 | -103.8402
25.1345 | -103.8338
25.1382 | -103.829
25.1419 | -103.8239
25.1456 | -103.8182
25.1493 | -103.8113
25.153 | -103.8053
25.1567 | -103.8001
25.1604 | -103.7942
25.1641 | -103.7892
25.1678 | -103.784
25.1715 | -103.779
25.1752 | -103.7725
25.1789 | -103.7655
25.1826 | -103.761
25.1864 | -103.7575
25.1901 | -103.7546
25.1938 | -103.7498
25.1975 | -103.7451
25.2012 | -103.7429
25.2049 | -103.7387
25.2086 | -103.7374
25.2123 | -103.7364
25.216 | -103.7368
25.2197 | -103.7381
25.2234 | -103.7376
25.2271 | -103.739
25.2308 | -103.7422
25.2345 | -103.7446
25.2382 | -103.7459
25.2419 | -103.7476
25.2456 | -103.7504
25.2493 | -103.754
25.253 | -103.7551
25.2567 | -103.7569
25.2604 | -103.7586
25.2641 | -103.7591
25.2678 | -103.7617
25.2715 | -103.7686
25.2752 | -103.7746
25.2789 | -103.7809
25.2827 | -103.7857
25.2864 | -103.79
25.2901 | -103.7941
25.2938 | -103.7988
25.2975 | -103.8028
25.3012 | -103.8071
25.3049 | -103.81
25.3086 | -103.8141
25.3123 | -103.8189
25.316 | -103.8238
25.3197 | -103.8283
25.3234 | -103.8327
25.3271 | -103.8366
25.3308 | -103.8404
25.3345 | -103.8436
25.3382 | -103.8464
25.3419 | -103.8491
25.3456 | -103.8509
25.3493 | -103.8525
25.353 | -103.8535
25.3567 | -103.8552
25.3604 | -103.8571
25.3641 | -103.859
25.3678 | -103.8607
25.3715 | -103.8601
25.3752 | -103.8583
25.379 | -103.8566
25.3827 | -103.8555
25.3864 | -103.8537
25.3901 | -103.8513
25.3938 | -103.8499
25.3975 | -103.8473
25.4012 | -103.8431
25.4049 | -103.8399
25.4086 | -103.8374
25.4123 | -103.8332
25.416 | -103.8287
25.4197 | -103.825
25.4234 | -103.8231
25.4271 | -103.8207
25.4308 | -103.8177
25.4345 | -103.814
25.4382 | -103.8117
25.4419 | -103.8081
25.4456 | -103.8053
25.4493 | -103.8034
25.453 | -103.8026
25.4567 | -103.8018
25.4604 | -103.8007
25.4641 | -103.8007
25.4678 | -103.8013
25.4716 | -103.8028
25.4753 | -103.8042
25.479 | -103.8069
25.4827 | -103.8105
25.4864 | -103.815
25.4901 | -103.8215
25.4938 | -103.8273
25.4975 | -103.8337
25.5012 | -103.8416
25.5049 | -103.8505
25.5086 | -103.8614
25.5123 | -103.8729
25.516 | -103.8841
25.5197 | -103.8968
25.5234 | -103.9104
25.5271 | -103.9241
25.5308 | -103.9371
25.5345 | -103.953
25.5382 | -103.9693
25.5419 | -103.9861
25.5456 | -104.0049
25.5493 | -104.023
25.553 | -104.0427
25.5567 | -104.0622
25.5604 | -104.0814
25.5641 | -104.1039
25.5678 | -104.1242
25.5716 | -104.144
25.5753 | -104.1664
25.579 | -104.1886
25.5827 | -104.2106
25.5864 | -104.2318
25.5901 | -104.2538
25.5938 | -104.2761
25.5975 | -104.2998
25.6012 | -104.324
25.6049 | -104.347
25.6086 | -104.3686
25.6123 | -104.3909
25.616 | -104.4129
25.6197 | -104.435
25.6234 | -104.4587
25.6271 | -104.4813
25.6308 | -104.5013
25.6345 | -104.5208
25.6382 | -104.539
25.6419 | -104.5592
25.6456 | -104.5734
25.6493 | -104.592
25.653 | -104.6094
25.6567 | -104.6259
25.6604 | -104.643
25.6642 | -104.6575
25.6679 | -104.6734
25.6716 | -104.6891
25.6753 | -104.7028
25.679 | -104.7172
25.6827 | -104.731
25.6864 | -104.7437
25.6901 | -104.7568
25.6938 | -104.7703
25.6975 | -104.7832
25.7012 | -104.7936
25.7049 | -104.805
25.7086 | -104.8166
25.7123 | -104.8262
25.716 | -104.8362
25.7197 | -104.8443
25.7234 | -104.8521
25.7271 | -104.8585
25.7308 | -104.8659
25.7345 | -104.8743
25.7382 | -104.8822
25.7419 | -104.8913
25.7456 | -104.8991
25.7493 | -104.9048
25.753 | -104.9094
25.7567 | -104.915
25.7605 | -104.9202
25.7642 | -104.9252
25.7679 | -104.9285
25.7716 | -104.9333
25.7753 | -104.9403
25.779 | -104.9466
25.7827 | -104.951
25.7864 | -104.9534
25.7901 | -104.9568
25.7938 | -104.9618
25.7975 | -104.9664
25.8012 | -104.9707
25.8049 | -104.9759
25.8086 | -104.9794
25.8123 | -104.9823
25.816 | -104.9873
25.8197 | -104.9922
25.8234 | -104.9959
25.8271 | -104.9986
25.8308 | -105.0035
25.8345 | -105.007
25.8382 | -105.0097
25.8419 | -105.0147
25.8456 | -105.0188
25.8493 | -105.022
25.853 | -105.0264
25.8568 | -105.0319
25.8605 | -105.0357
25.8642 | -105.0382
25.8679 | -105.0407
25.8716 | -105.0433
25.8753 | -105.0462
25.879 | -105.0507
25.8827 | -105.0539
25.8864 | -105.0578
25.8901 | -105.0621
25.8938 | -105.067
25.8975 | -105.0727
25.9012 | -105.0789
25.9049 | -105.0826
25.9086 | -105.0851
25.9123 | -105.0896
25.916 | -105.0935
25.9197 | -105.0969
25.9234 | -105.0996
25.9271 | -105.1028
25.9308 | -105.1065
25.9345 | -105.1124
25.9382 | -105.118
25.9419 | -105.124
25.9456 | -105.13
25.9493 | -105.1373
25.9531 | -105.1436
25.9568 | -105.1499
25.9605 | -105.155
25.9642 | -105.162
25.9679 | -105.171
25.9716 | -105.18
25.9753 | -105.1905
25.979 | -105.1986
25.9827 | -105.2053
25.9864 | -105.2103
25.9901 | -105.2176
25.9938 | -105.2254
25.9975 | -105.2319
26.0012 | -105.2385
26.0049 | -105.2457
26.0086 | -105.2525
26.0123 | -105.2606
26.016 | -105.2687
26.0197 | -105.275
26.0234 | -105.2813
26.0271 | -105.2863
26.0308 | -105.2919
26.0345 | -105.2977
26.0382 | -105.3037
26.0419 | -105.3099
26.0456 | -105.3152
26.0494 | -105.3196
26.0531 | -105.3229
26.0568 | -105.3264
26.0605 | -105.3281
26.0642 | -105.3308
26.0679 | -105.3362
26.0716 | -105.3391
26.0753 | -105.3415
26.079 | -105.3442
26.0827 | -105.3469
26.0864 | -105.3502
26.0901 | -105.3528
26.0938 | -105.3558
26.0975 | -105.3585
26.1012 | -105.3617
26.1049 | -105.3648
26.1086 | -105.3653
26.1123 | -105.3659
26.116 | -105.3689
26.1197 | -105.3729
26.1234 | -105.3759
26.1271 | -105.378
26.1308 | -105.3799
26.1345 | -105.3835
26.1382 | -105.3868
26.142 | -105.3884
26.1457 | -105.391
26.1494 | -105.3932
26.1531 | -105.3964
26.1568 | -105.3996
26.1605 | -105.4008
26.1642 | -105.4045
26.1679 | -105.4069
26.1716 | -105.4092
26.1753 | -105.4129
26.179 | -105.4179
26.1827 | -105.423
26.1864 | -105.4268
26.1901 | -105.4321
26.1938 | -105.4375
26.1975 | -105.4437
26.2012 | -105.4455
26.2049 | -105.448
26.2086 | -105.4521
26.2123 | -105.4559
26.216 | -105.4593
26.2197 | -105.4639
26.2234 | -105.4688
26.2271 | -105.4748
26.2308 | -105.4813
26.2345 | -105.4888
26.2383 | -105.4961
26.242 | -105.5028
26.2457 | -105.5074
26.2494 | -105.5115
26.2531 | -105.5165
26.2568 | -105.5212
26.2605 | -105.5259
26.2642 | -105.531
26.2679 | -105.5331
26.2716 | -105.5355
26.2753 | -105.5401
26.279 | -105.5445
26.2827 | -105.5496
26.2864 | -105.5561
26.2901 | -105.5626
26.2938 | -105.569
26.2975 | -105.5746
26.3012 | -105.5811
26.3049 | -105.5861
26.3086 | -105.5902
26.3123 | -105.595
26.316 | -105.6009
26.3197 | -105.6064
26.3234 | -105.6107
26.3271 | -105.6164
26.3308 | -105.6211
26.3346 | -105.6258
26.3383 | -105.6312
26.342 | -105.635
26.3457 | -105.6382
26.3494 | -105.644
26.3531 | -105.6507
26.3568 | -105.6572
26.3605 | -105.6618
26.3642 | -105.6664
26.3679 | -105.6714
26.3716 | -105.6775
26.3753 | -105.6816
26.379 | -105.6855
26.3827 | -105.6879
26.3864 | -105.6912
26.3901 | -105.6971
26.3938 | -105.7023
26.3975 | -105.707
26.4012 | -105.7104
26.4049 | -105.7136
26.4086 | -105.7168
26.4123 | -105.7203
26.416 | -105.724
26.4197 | -105.7268
26.4234 | -105.7302
26.4271 | -105.7327
26.4309 | -105.7365
26.4346 | -105.7397
26.4383 | -105.7417
26.442 | -105.744
26.4457 | -105.746
26.4494 | -105.749
26.4531 | -105.7496
26.4568 | -105.7481
26.4605 | -105.7492
26.4642 | -105.7497
26.4679 | -105.749
26.4716 | -105.7475
26.4753 | -105.7482
26.479 | -105.7498
26.4827 | -105.751
26.4864 | -105.7504
26.4901 | -105.7504
26.4938 | -105.7507
26.4975 | -105.7497
26.5012 | -105.751
26.5049 | -105.7502
26.5086 | -105.7484
26.5123 | -105.7485
26.516 | -105.7474
26.5197 | -105.747
26.5235 | -105.7465
26.5272 | -105.7444
26.5309 | -105.7438
26.5346 | -105.7419
26.5383 | -105.7401
26.542 | -105.7376
26.5457 | -105.7369
26.5494 | -105.7349
26.5531 | -105.7334
26.5568 | -105.7319
26.5605 | -105.7301
26.5642 | -105.7276
26.5679 | -105.7243
26.5716 | -105.7213
26.5753 | -105.7216
26.579 | -105.7224
26.5827 | -105.7219
26.5864 | -105.7203
26.5901 | -105.7209
26.5938 | -105.7204
26.5975 | -105.7168
26.6012 | -105.7146
26.6049 | -105.7121
26.6086 | -105.7083
26.6123 | -105.7064
26.616 | -105.7045
26.6198 | -105.7036
26.6235 | -105.7034
26.6272 | -105.704
26.6309 | -105.7037
26.6346 | -105.7037
26.6383 | -105.7051
26.642 | -105.7072
26.6457 | -105.7094
26.6494 | -105.7108
26.6531 | -105.711
26.6568 | -105.7095
26.6605 | -105.7091
26.6642 | -105.7089
26.6679 | -105.7084
26.6716 | -105.7085
26.6753 | -105.7074
26.679 | -105.7059
26.6827 | -105.7054
26.6864 | -105.7056
26.6901 | -105.7067
26.6938 | -105.7069
26.6975 | -105.708
26.7012 | -105.7094
26.7049 | -105.7116
26.7086 | -105.7111
26.7123 | -105.709
26.7161 | -105.7066
26.7198 | -105.706
26.7235 | -105.706
26.7272 | -105.7041
26.7309 | -105.7029
26.7346 | -105.7029
26.7383 | -105.7016
26.742 | -105.7006
26.7457 | -105.7005
26.7494 | -105.7001
26.7531 | -105.6998
26.7568 | -105.698
26.7605 | -105.6989
26.7642 | -105.6981
26.7679 | -105.696
26.7716 | -105.6947
26.7753 | -105.6947
26.779 | -105.6951
26.7827 | -105.6951
26.7864 | -105.6942
26.7901 | -105.6939
26.7938 | -105.694
26.7975 | -105.696
26.8012 | -105.6968
26.8049 | -105.6994
26.8086 | -105.7015
26.8124 | -105.7021
26.8161 | -105.7028
26.8198 | -105.7061
26.8235 | -105.7083
26.8272 | -105.7086
26.8309 | -105.7092
26.8346 | -105.711
26.8383 | -105.7118
26.842 | -105.7133
26.8457 | -105.7151
26.8494 | -105.7176
26.8531 | -105.7175
26.8568 | -105.7178
26.8605 | -105.7191
26.8642 | -105.7215
26.8679 | -105.7242
26.8716 | -105.7265
26.8753 | -105.7302
26.879 | -105.7301
26.8827 | -105.7315
26.8864 | -105.7336
26.8901 | -105.7338
26.8938 | -105.7349
26.8975 | -105.7374
26.9012 | -105.7387
26.9049 | -105.7404
26.9087 | -105.742
26.9124 | -105.7428
26.9161 | -105.7448
26.9198 | -105.7477
26.9235 | -105.7495
26.9272 | -105.7512
26.9309 | -105.753
26.9346 | -105.7553
26.9383 | -105.7566
26.942 | -105.7582
26.9457 | -105.7618
26.9494 | -105.7657
26.9531 | -105.7685
26.9568 | -105.7706
26.9605 | -105.7733
26.9642 | -105.7754
26.9679 | -105.7767
26.9716 | -105.7795
26.9753 | -105.7838
26.979 | -105.7858
26.9827 | -105.7888
26.9864 | -105.7932
26.9901 | -105.7974
26.9938 | -105.802
26.9975 | -105.8062
27.0012 | -105.8096
27.005 | -105.8126
27.0087 | -105.8164
27.0124 | -105.8209
27.0161 | -105.8245
27.0198 | -105.8281
27.0235 | -105.8314
27.0272 | -105.835
27.0309 | -105.8387
27.0346 | -105.8446
27.0383 | -105.8497
27.042 | -105.8558
27.0457 | -105.8621
27.0494 | -105.8675
27.0531 | -105.8743
27.0568 | -105.8816
27.0605 | -105.8878
27.0642 | -105.8931
27.0679 | -105.8969
27.0716 | -105.9024
27.0753 | -105.9073
27.079 | -105.9113
27.0827 | -105.916
27.0864 | -105.9209
27.0901 | -105.9267
27.0938 | -105.9331
27.0976 | -105.9393
27.1013 | -105.9435
27.105 | -105.9463
27.1087 | -105.9507
27.1124 | -105.9538
27.1161 | -105.955
27.1198 | -105.9576
27.1235 | -105.9595
27.1272 | -105.9604
27.1309 | -105.9601
27.1346 | -105.959
27.1383 | -105.9589
27.142 | -105.9575
27.1457 | -105.9573
27.1494 | -105.9563
27.1531 | -105.9533
27.1568 | -105.9492
27.1605 | -105.9462
27.1642 | -105.9426
27.1679 | -105.9379
27.1716 | -105.9323
27.1753 | -105.9272
27.179 | -105.9214
27.1827 | -105.9161
27.1864 | -105.9112
27.1901 | -105.9063
27.1939 | -105.9016
27.1976 | -105.895
27.2013 | -105.8881
27.205 | -105.8822
27.2087 | -105.8784
27.2124 | -105.8733
27.2161 | -105.8671
27.2198 | -105.8598
27.2235 | -105.8528
27.2272 | -105.8444
27.2309 | -105.8372
27.2346 | -105.8303
27.2383 | -105.8238
27.242 | -105.8172
27.2457 | -105.8106
27.2494 | -105.802
27.2531 | -105.7946
27.2568 | -105.7857
27.2605 | -105.7763
27.2642 | -105.771
27.2679 | -105.764
27.2716 | -105.7576
27.2753 | -105.7517
27.279 | -105.7439
27.2827 | -105.7352
27.2864 | -105.7275
27.2902 | -105.721
27.2939 | -105.7144
27.2976 | -105.7081
27.3013 | -105.701
27.305 | -105.693
27.3087 | -105.6838
27.3124 | -105.6759
27.3161 | -105.6655
27.3198 | -105.6561
27.3235 | -105.6487
27.3272 | -105.6412
27.3309 | -105.6324
27.3346 | -105.6238
27.3383 | -105.6148
27.342 | -105.6053
27.3457 | -105.595
27.3494 | -105.5843
27.3531 | -105.5739
27.3568 | -105.5639
27.3605 | -105.5538
27.3642 | -105.5432
27.3679 | -105.5332
27.3716 | -105.5224
27.3753 | -105.5086
27.379 | -105.496
27.3827 | -105.4855
27.3865 | -105.4725
27.3902 | -105.4602
27.3939 | -105.4486
27.3976 | -105.4347
27.4013 | -105.4209
27.405 | -105.4078
27.4087 | -105.3958
27.4124 | -105.3833
27.4161 | -105.3722
27.4198 | -105.3595
27.4235 | -105.3455
27.4272 | -105.331
27.4309 | -105.3162
27.4346 | -105.3024
27.4383 | -105.2884
27.442 | -105.2752
27.4457 | -105.2621
27.4494 | -105.2484
27.4531 | -105.2363
27.4568 | -105.2221
27.4605 | -105.2073
27.4642 | -105.1935
27.4679 | -105.1793
27.4716 | -105.1664
27.4753 | -105.1533
27.4791 | -105.1373
27.4828 | -105.123
27.4865 | -105.1083
27.4902 | -105.0934
27.4939 | -105.0783
27.4976 | -105.0627
27.5013 | -105.0483
27.505 | -105.0344
27.5087 | -105.0207
27.5124 | -105.0082
27.5161 | -104.9963
27.5198 | -104.9851
27.5235 | -104.9735
27.5272 | -104.962
27.5309 | -104.9499
27.5346 | -104.9381
27.5383 | -104.9268
27.542 | -104.9178
27.5457 | -104.909
27.5494 | -104.9
27.5531 | -104.8904
27.5568 | -104.8818
27.5605 | -104.8726
27.5642 | -104.8651
27.5679 | -104.8565
27.5716 | -104.8486
27.5754 | -104.8403
27.5791 | -104.8333
27.5828 | -104.827
27.5865 | -104.8214
27.5902 | -104.8159
27.5939 | -104.8102
27.5976 | -104.8055
27.6013 | -104.8011
27.605 | -104.7983
27.6087 | -104.7936
27.6124 | -104.7897
27.6161 | -104.7859
27.6198 | -104.7824
27.6235 | -104.7778
27.6272 | -104.7735
27.6309 | -104.7681
27.6346 | -104.7658
27.6383 | -104.7643
27.642 | -104.7605
27.6457 | -104.7585
27.6494 | -104.7568
27.6531 | -104.7539
27.6568 | -104.7516
27.6605 | -104.7472
27.6642 | -104.7438
27.6679 | -104.744
27.6717 | -104.7442
27.6754 | -104.7436
27.6791 | -104.7417
27.6828 | -104.7383
27.6865 | -104.7374
27.6902 | -104.7372
27.6939 | -104.7381
27.6976 | -104.7397
27.7013 | -104.7392
27.705 | -104.7391
27.7087 | -104.7373
27.7124 | -104.7365
27.7161 | -104.7339
27.7198 | -104.7323
27.7235 | -104.7291
27.7272 | -104.7242
27.7309 | -104.7225
27.7346 | -104.7223
27.7383 | -104.7223
27.742 | -104.7208
27.7457 | -104.7183
27.7494 | -104.7166
27.7531 | -104.7153
27.7568 | -104.7115
27.7605 | -104.7082
27.7642 | -104.705
27.768 | -104.7007
27.7717 | -104.6967
27.7754 | -104.6943
27.7791 | -104.6895
27.7828 | -104.6831
27.7865 | -104.6765
27.7902 | -104.6697
27.7939 | -104.662
27.7976 | -104.6558
27.8013 | -104.6473
27.805 | -104.6391
27.8087 | -104.6303
27.8124 | -104.6218
27.8161 | -104.615
27.8198 | -104.6068
27.8235 | -104.598
27.8272 | -104.5881
27.8309 | -104.5777
27.8346 | -104.5668
27.8383 | -104.5569
27.842 | -104.5449
27.8457 | -104.5353
27.8494 | -104.5255
27.8531 | -104.5168
27.8568 | -104.5061
27.8605 | -104.4951
27.8643 | -104.4831
27.868 | -104.4703
27.8717 | -104.4606
27.8754 | -104.4488
27.8791 | -104.4369
27.8828 | -104.4252
27.8865 | -104.4139
27.8902 | -104.4038
27.8939 | -104.3923
27.8976 | -104.3798
27.9013 | -104.3677
27.905 | -104.3552
27.9087 | -104.3429
27.9124 | -104.3307
27.9161 | -104.3184
27.9198 | -104.3068
27.9235 | -104.296
27.9272 | -104.2847
27.9309 | -104.2736
27.9346 | -104.2613
27.9383 | -104.2491
27.942 | -104.2387
27.9457 | -104.2281
27.9494 | -104.218
27.9531 | -104.2091
27.9568 | -104.2009
27.9606 | -104.1944
27.9643 | -104.1862
27.968 | -104.1787
27.9717 | -104.1712
27.9754 | -104.1633
27.9791 | -104.1554
27.9828 | -104.1511
27.9865 | -104.1473
27.9902 | -104.1436
27.9939 | -104.1416
27.9976 | -104.1391
28.0013 | -104.137