## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 240 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 254
Hood | 02-92B/P | 181
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1001 | hPa |  | PRC PRO 01
Humidity | 59.1 | %/Hr
Internal temperature of calibration bench | 23.17 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1011 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.4 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.99 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.02 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3112 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 31/05/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 05/06/2018
Technical Manager | Anthony HUE | Approval |  | 13/06/2018
CEO | Robert GLEMOT | Approval |  | 13/06/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4471 | 57.4669
0.012209 | -36.0198 | 38.2781
0.018313 | -34.9588 | 28.1603
0.024417 | -34.4978 | 21.9379
0.030521 | -34.3875 | 18.1829
0.036626 | -34.2428 | 15.1347
0.04273 | -34.1466 | 12.7692
0.048834 | -34.0997 | 11.2484
0.054938 | -34.0341 | 10.1127
0.061043 | -33.9737 | 9.0332
0.067147 | -33.9637 | 8.2156
0.073251 | -33.9651 | 7.4608
0.079355 | -33.9842 | 6.8047
0.08546 | -33.9989 | 6.2144
0.091564 | -33.9942 | 5.7764
0.097668 | -33.9754 | 5.5293
0.10377 | -33.9521 | 5.2799
0.10988 | -33.934 | 5.0017
0.11598 | -33.9245 | 4.7084
0.12209 | -33.9267 | 4.425
0.12819 | -33.9371 | 4.1419
0.13429 | -33.9483 | 3.8536
0.1404 | -33.9518 | 3.5743
0.1465 | -33.954 | 3.3161
0.15261 | -33.9516 | 3.096
0.15871 | -33.9452 | 2.9343
0.16482 | -33.9371 | 2.8326
0.17092 | -33.9345 | 2.7941
0.17702 | -33.9321 | 2.7221
0.18313 | -33.931 | 2.644
0.18923 | -33.9303 | 2.5347
0.19534 | -33.9288 | 2.4076
0.20144 | -33.9274 | 2.2653
0.20754 | -33.9246 | 2.1458
0.21365 | -33.9236 | 2.0323
0.21975 | -33.9232 | 1.936
0.22586 | -33.9243 | 1.8504
0.23196 | -33.9268 | 1.7796
0.23807 | -33.931 | 1.7183
0.24417 | -33.935 | 1.6604
0.25027 | -33.939 | 1.599
0.25638 | -33.9414 | 1.5379
0.26248 | -33.9432 | 1.472
0.26859 | -33.944 | 1.4073
0.27469 | -33.9443 | 1.3365
0.2808 | -33.9451 | 1.2627
0.2869 | -33.9458 | 1.1917
0.293 | -33.9466 | 1.1244
0.29911 | -33.9474 | 1.0641
0.30521 | -33.9482 | 1.0049
0.31132 | -33.9488 | 0.96613
0.31742 | -33.95 | 0.92323
0.32353 | -33.9503 | 0.88838
0.32963 | -33.9511 | 0.8595
0.33573 | -33.9521 | 0.83963
0.34184 | -33.9535 | 0.81876
0.34794 | -33.9557 | 0.79583
0.35405 | -33.9579 | 0.77251
0.36015 | -33.9605 | 0.74801
0.36626 | -33.9626 | 0.72138
0.37236 | -33.9648 | 0.68554
0.37846 | -33.967 | 0.64197
0.38457 | -33.9679 | 0.58514
0.39067 | -33.969 | 0.5389
0.39678 | -33.9696 | 0.48734
0.40288 | -33.9701 | 0.43341
0.40899 | -33.9699 | 0.38154
0.41509 | -33.9691 | 0.33059
0.42119 | -33.969 | 0.28708
0.4273 | -33.9688 | 0.24072
0.4334 | -33.9686 | 0.1944
0.43951 | -33.9688 | 0.15093
0.44561 | -33.9695 | 0.11105
0.45172 | -33.9711 | 0.079377
0.45782 | -33.9718 | 0.042281
0.46392 | -33.973 | 0.0060158
0.47003 | -33.974 | -0.026303
0.47613 | -33.9748 | -0.060361
0.48224 | -33.9753 | -0.095114
0.48834 | -33.9743 | -0.12976
0.49445 | -33.974 | -0.16497
0.50055 | -33.9734 | -0.19697
0.50665 | -33.9726 | -0.23109
0.51276 | -33.9716 | -0.26509
0.51886 | -33.9705 | -0.29902
0.52497 | -33.9692 | -0.32307
0.53107 | -33.9684 | -0.35247
0.53717 | -33.9677 | -0.38061
0.54328 | -33.9668 | -0.40332
0.54938 | -33.9666 | -0.42565
0.55549 | -33.9666 | -0.44646
0.56159 | -33.9677 | -0.45794
0.5677 | -33.9681 | -0.4724
0.5738 | -33.9691 | -0.48859
0.5799 | -33.97 | -0.50346
0.58601 | -33.9711 | -0.51347
0.59211 | -33.9728 | -0.5196
0.59822 | -33.9742 | -0.52894
0.60432 | -33.9757 | -0.54091
0.61043 | -33.977 | -0.54872
0.61653 | -33.9783 | -0.55886
0.62263 | -33.9799 | -0.56806
0.62874 | -33.9806 | -0.57853
0.63484 | -33.9814 | -0.58762
0.64095 | -33.9819 | -0.59768
0.64705 | -33.9823 | -0.60723
0.65316 | -33.9825 | -0.61947
0.65926 | -33.9826 | -0.6351
0.66536 | -33.9814 | -0.65228
0.67147 | -33.981 | -0.66527
0.67757 | -33.9806 | -0.67933
0.68368 | -33.9798 | -0.69804
0.68978 | -33.979 | -0.71508
0.69589 | -33.978 | -0.73392
0.70199 | -33.977 | -0.7566
0.70809 | -33.9761 | -0.77565
0.7142 | -33.9753 | -0.79603
0.7203 | -33.9741 | -0.81511
0.72641 | -33.9734 | -0.83752
0.73251 | -33.9727 | -0.85746
0.73862 | -33.9729 | -0.87928
0.74472 | -33.9725 | -0.89922
0.75082 | -33.9722 | -0.92133
0.75693 | -33.9725 | -0.94317
0.76303 | -33.973 | -0.96521
0.76914 | -33.9749 | -0.98576
0.77524 | -33.9758 | -1.0054
0.78135 | -33.9771 | -1.0255
0.78745 | -33.9785 | -1.0455
0.79355 | -33.9797 | -1.0653
0.79966 | -33.9815 | -1.0848
0.80576 | -33.9838 | -1.1045
0.81187 | -33.9854 | -1.1261
0.81797 | -33.9874 | -1.1474
0.82408 | -33.9891 | -1.168
0.83018 | -33.9908 | -1.1886
0.83628 | -33.9927 | -1.2071
0.84239 | -33.9939 | -1.2276
0.84849 | -33.9954 | -1.2494
0.8546 | -33.9968 | -1.2692
0.8607 | -33.9981 | -1.29
0.86681 | -33.9992 | -1.309
0.87291 | -34.0002 | -1.3278
0.87901 | -34.0006 | -1.3458
0.88512 | -34.0014 | -1.3639
0.89122 | -34.0021 | -1.3821
0.89733 | -34.0025 | -1.3977
0.90343 | -34.0029 | -1.4152
0.90953 | -34.003 | -1.4295
0.91564 | -34.003 | -1.4453
0.92174 | -34.0032 | -1.4602
0.92785 | -34.0033 | -1.4747
0.93395 | -34.0034 | -1.4884
0.94006 | -34.0035 | -1.5011
0.94616 | -34.0032 | -1.5128
0.95226 | -34.0033 | -1.5253
0.95837 | -34.0034 | -1.5391
0.96447 | -34.0035 | -1.5517
0.97058 | -34.0033 | -1.565
0.97668 | -34.0031 | -1.5782
0.98279 | -34.0028 | -1.5908
0.98889 | -34.0026 | -1.6013
0.99499 | -34.0025 | -1.6145
1.0011 | -34.0024 | -1.625
1.0072 | -34.002 | -1.6364
1.0133 | -34.0016 | -1.6469
1.0194 | -34.0012 | -1.6596
1.0255 | -34.0008 | -1.6732
1.0316 | -34.0002 | -1.6849
1.0377 | -33.9998 | -1.6984
1.0438 | -33.9992 | -1.7109
1.0499 | -33.9984 | -1.726
1.056 | -33.9977 | -1.7401
1.0621 | -33.997 | -1.7542
1.0682 | -33.9965 | -1.7664
1.0743 | -33.9958 | -1.7817
1.0805 | -33.9953 | -1.7959
1.0866 | -33.9944 | -1.8108
1.0927 | -33.9935 | -1.8251
1.0988 | -33.9928 | -1.8398
1.1049 | -33.992 | -1.8537
1.111 | -33.9912 | -1.8689
1.1171 | -33.9905 | -1.8832
1.1232 | -33.99 | -1.8982
1.1293 | -33.9895 | -1.9121
1.1354 | -33.9892 | -1.9269
1.1415 | -33.9888 | -1.9422
1.1476 | -33.9882 | -1.9575
1.1537 | -33.9877 | -1.9721
1.1598 | -33.9878 | -1.9866
1.1659 | -33.9875 | -1.9997
1.172 | -33.9873 | -2.0134
1.1781 | -33.987 | -2.0272
1.1842 | -33.9869 | -2.0412
1.1903 | -33.9872 | -2.052
1.1964 | -33.9871 | -2.0652
1.2025 | -33.9872 | -2.079
1.2086 | -33.9876 | -2.0912
1.2147 | -33.9876 | -2.1035
1.2209 | -33.9879 | -2.1145
1.227 | -33.9887 | -2.1251
1.2331 | -33.9891 | -2.1369
1.2392 | -33.9895 | -2.1485
1.2453 | -33.9899 | -2.1594
1.2514 | -33.9903 | -2.1691
1.2575 | -33.9909 | -2.1799
1.2636 | -33.9918 | -2.1887
1.2697 | -33.9924 | -2.1977
1.2758 | -33.9931 | -2.2087
1.2819 | -33.9938 | -2.2185
1.288 | -33.9945 | -2.2291
1.2941 | -33.9953 | -2.2374
1.3002 | -33.996 | -2.2452
1.3063 | -33.9965 | -2.2541
1.3124 | -33.9972 | -2.263
1.3185 | -33.998 | -2.2718
1.3246 | -33.9986 | -2.2808
1.3307 | -33.9995 | -2.2892
1.3368 | -34.0003 | -2.2977
1.3429 | -34.0009 | -2.3065
1.349 | -34.0014 | -2.3156
1.3551 | -34.0021 | -2.3248
1.3613 | -34.0027 | -2.3329
1.3674 | -34.0033 | -2.3428
1.3735 | -34.0038 | -2.3524
1.3796 | -34.0042 | -2.3613
1.3857 | -34.0045 | -2.3703
1.3918 | -34.0051 | -2.3803
1.3979 | -34.0056 | -2.3893
1.404 | -34.0061 | -2.4004
1.4101 | -34.0065 | -2.4106
1.4162 | -34.0069 | -2.421
1.4223 | -34.0074 | -2.4322
1.4284 | -34.0078 | -2.442
1.4345 | -34.0081 | -2.4534
1.4406 | -34.0084 | -2.4666
1.4467 | -34.0088 | -2.4781
1.4528 | -34.0093 | -2.4893
1.4589 | -34.0097 | -2.5009
1.465 | -34.01 | -2.5116
1.4711 | -34.0105 | -2.5227
1.4772 | -34.011 | -2.5359
1.4833 | -34.0114 | -2.5487
1.4894 | -34.0119 | -2.5607
1.4955 | -34.0125 | -2.5725
1.5016 | -34.013 | -2.5848
1.5078 | -34.0134 | -2.5985
1.5139 | -34.0138 | -2.6113
1.52 | -34.0142 | -2.6237
1.5261 | -34.0145 | -2.6361
1.5322 | -34.0148 | -2.65
1.5383 | -34.0152 | -2.6624
1.5444 | -34.0157 | -2.6748
1.5505 | -34.0162 | -2.688
1.5566 | -34.0167 | -2.7008
1.5627 | -34.0172 | -2.7144
1.5688 | -34.0178 | -2.7278
1.5749 | -34.0181 | -2.7402
1.581 | -34.0184 | -2.7533
1.5871 | -34.0188 | -2.7668
1.5932 | -34.0192 | -2.7795
1.5993 | -34.0196 | -2.7935
1.6054 | -34.0199 | -2.8065
1.6115 | -34.0204 | -2.8205
1.6176 | -34.0206 | -2.8328
1.6237 | -34.0208 | -2.8462
1.6298 | -34.0211 | -2.8585
1.6359 | -34.0215 | -2.8711
1.642 | -34.0219 | -2.8827
1.6482 | -34.0221 | -2.8955
1.6543 | -34.0224 | -2.909
1.6604 | -34.0227 | -2.9207
1.6665 | -34.0228 | -2.933
1.6726 | -34.023 | -2.9453
1.6787 | -34.0233 | -2.9576
1.6848 | -34.0232 | -2.9699
1.6909 | -34.0234 | -2.9824
1.697 | -34.0236 | -2.9944
1.7031 | -34.0237 | -3.0067
1.7092 | -34.0236 | -3.0187
1.7153 | -34.0238 | -3.0304
1.7214 | -34.0235 | -3.0424
1.7275 | -34.0235 | -3.0548
1.7336 | -34.0234 | -3.0676
1.7397 | -34.0235 | -3.0797
1.7458 | -34.0235 | -3.0917
1.7519 | -34.0235 | -3.1037
1.758 | -34.0235 | -3.1154
1.7641 | -34.0236 | -3.1269
1.7702 | -34.0235 | -3.1396
1.7763 | -34.0235 | -3.1509
1.7824 | -34.0234 | -3.1632
1.7885 | -34.0234 | -3.1755
1.7947 | -34.0235 | -3.1873
1.8008 | -34.0234 | -3.1984
1.8069 | -34.0234 | -3.2098
1.813 | -34.0235 | -3.2218
1.8191 | -34.0234 | -3.2344
1.8252 | -34.0234 | -3.247
1.8313 | -34.0236 | -3.2591
1.8374 | -34.0237 | -3.2708
1.8435 | -34.0237 | -3.2826
1.8496 | -34.0237 | -3.2955
1.8557 | -34.0239 | -3.3082
1.8618 | -34.0241 | -3.3203
1.8679 | -34.0242 | -3.3322
1.874 | -34.0243 | -3.3441
1.8801 | -34.0244 | -3.3564
1.8862 | -34.0246 | -3.3682
1.8923 | -34.0247 | -3.3803
1.8984 | -34.0248 | -3.3922
1.9045 | -34.0249 | -3.4037
1.9106 | -34.0249 | -3.4158
1.9167 | -34.0249 | -3.4279
1.9228 | -34.0251 | -3.4397
1.9289 | -34.0252 | -3.4515
1.9351 | -34.0253 | -3.4622
1.9412 | -34.0255 | -3.4733
1.9473 | -34.0258 | -3.4853
1.9534 | -34.0259 | -3.4967
1.9595 | -34.0261 | -3.5084
1.9656 | -34.0264 | -3.5206
1.9717 | -34.0266 | -3.5322
1.9778 | -34.0269 | -3.5435
1.9839 | -34.0271 | -3.5549
1.99 | -34.0273 | -3.5659
1.9961 | -34.0274 | -3.5779
2.0022 | -34.0276 | -3.5892
2.0083 | -34.0277 | -3.6004
2.0144 | -34.0279 | -3.6116
2.0205 | -34.0282 | -3.6233
2.0266 | -34.0285 | -3.6346
2.0327 | -34.0287 | -3.6461
2.0388 | -34.0291 | -3.6576
2.0449 | -34.0293 | -3.6689
2.051 | -34.0296 | -3.6801
2.0571 | -34.0299 | -3.6914
2.0632 | -34.0301 | -3.7027
2.0693 | -34.0303 | -3.7148
2.0754 | -34.0305 | -3.726
2.0816 | -34.0308 | -3.738
2.0877 | -34.031 | -3.7494
2.0938 | -34.0314 | -3.7615
2.0999 | -34.0316 | -3.773
2.106 | -34.0318 | -3.7845
2.1121 | -34.0319 | -3.7964
2.1182 | -34.0322 | -3.807
2.1243 | -34.0325 | -3.8176
2.1304 | -34.0327 | -3.8287
2.1365 | -34.033 | -3.8403
2.1426 | -34.0332 | -3.8521
2.1487 | -34.0335 | -3.8632
2.1548 | -34.0339 | -3.8749
2.1609 | -34.0343 | -3.8862
2.167 | -34.0347 | -3.8973
2.1731 | -34.035 | -3.9092
2.1792 | -34.0352 | -3.9201
2.1853 | -34.0356 | -3.9319
2.1914 | -34.0359 | -3.9432
2.1975 | -34.0363 | -3.9546
2.2036 | -34.0365 | -3.9666
2.2097 | -34.0368 | -3.9779
2.2158 | -34.0371 | -3.9884
2.222 | -34.0374 | -3.9987
2.2281 | -34.0376 | -4.01
2.2342 | -34.038 | -4.0213
2.2403 | -34.0381 | -4.0317
2.2464 | -34.0384 | -4.0423
2.2525 | -34.0387 | -4.0536
2.2586 | -34.039 | -4.064
2.2647 | -34.0393 | -4.0743
2.2708 | -34.0396 | -4.0854
2.2769 | -34.04 | -4.0959
2.283 | -34.0402 | -4.107
2.2891 | -34.0404 | -4.1177
2.2952 | -34.0406 | -4.1281
2.3013 | -34.0408 | -4.1389
2.3074 | -34.0411 | -4.1491
2.3135 | -34.0413 | -4.1604
2.3196 | -34.0415 | -4.1704
2.3257 | -34.0417 | -4.181
2.3318 | -34.0418 | -4.192
2.3379 | -34.042 | -4.2026
2.344 | -34.0421 | -4.2139
2.3501 | -34.0423 | -4.2244
2.3562 | -34.0425 | -4.2347
2.3623 | -34.0426 | -4.2449
2.3685 | -34.0427 | -4.2553
2.3746 | -34.0429 | -4.2664
2.3807 | -34.0431 | -4.2767
2.3868 | -34.0433 | -4.2873
2.3929 | -34.0434 | -4.2978
2.399 | -34.0435 | -4.3082
2.4051 | -34.0438 | -4.3186
2.4112 | -34.044 | -4.3291
2.4173 | -34.0441 | -4.3396
2.4234 | -34.0442 | -4.3503
2.4295 | -34.0444 | -4.3609
2.4356 | -34.0445 | -4.3717
2.4417 | -34.0446 | -4.3827
2.4478 | -34.0447 | -4.3934
2.4539 | -34.0449 | -4.4035
2.46 | -34.045 | -4.4148
2.4661 | -34.0451 | -4.4257
2.4722 | -34.0453 | -4.4358
2.4783 | -34.0454 | -4.446
2.4844 | -34.0454 | -4.4569
2.4905 | -34.0455 | -4.4678
2.4966 | -34.0457 | -4.4784
2.5027 | -34.0458 | -4.4888
2.5089 | -34.0459 | -4.4995
2.515 | -34.0459 | -4.5102
2.5211 | -34.046 | -4.5208
2.5272 | -34.046 | -4.5316
2.5333 | -34.0462 | -4.5424
2.5394 | -34.0463 | -4.553
2.5455 | -34.0464 | -4.5643
2.5516 | -34.0465 | -4.5745
2.5577 | -34.0467 | -4.5858
2.5638 | -34.0469 | -4.5974
2.5699 | -34.0469 | -4.6086
2.576 | -34.0469 | -4.6201
2.5821 | -34.047 | -4.6312
2.5882 | -34.0471 | -4.6422
2.5943 | -34.0472 | -4.6529
2.6004 | -34.0474 | -4.6637
2.6065 | -34.0475 | -4.674
2.6126 | -34.0476 | -4.6843
2.6187 | -34.0477 | -4.6946
2.6248 | -34.0478 | -4.7052
2.6309 | -34.048 | -4.7165
2.637 | -34.0482 | -4.728
2.6431 | -34.0484 | -4.7383
2.6492 | -34.0484 | -4.7495
2.6554 | -34.0486 | -4.7602
2.6615 | -34.0487 | -4.771
2.6676 | -34.049 | -4.7823
2.6737 | -34.0492 | -4.7933
2.6798 | -34.0495 | -4.8043
2.6859 | -34.0496 | -4.8154
2.692 | -34.0498 | -4.8261
2.6981 | -34.05 | -4.8371
2.7042 | -34.0501 | -4.8481
2.7103 | -34.0503 | -4.8586
2.7164 | -34.0506 | -4.8694
2.7225 | -34.0509 | -4.8806
2.7286 | -34.0511 | -4.8915
2.7347 | -34.0512 | -4.9019
2.7408 | -34.0515 | -4.9118
2.7469 | -34.0517 | -4.923
2.753 | -34.052 | -4.9334
2.7591 | -34.0523 | -4.9446
2.7652 | -34.0525 | -4.9559
2.7713 | -34.0526 | -4.9665
2.7774 | -34.0529 | -4.9773
2.7835 | -34.053 | -4.9881
2.7896 | -34.0532 | -4.9985
2.7958 | -34.0535 | -5.0091
2.8019 | -34.0538 | -5.0193
2.808 | -34.054 | -5.0299
2.8141 | -34.0543 | -5.0409
2.8202 | -34.0546 | -5.0516
2.8263 | -34.0549 | -5.0624
2.8324 | -34.0552 | -5.0736
2.8385 | -34.0556 | -5.0843
2.8446 | -34.056 | -5.0949
2.8507 | -34.0563 | -5.1056
2.8568 | -34.0567 | -5.1159
2.8629 | -34.0569 | -5.1263
2.869 | -34.0572 | -5.1373
2.8751 | -34.0575 | -5.1479
2.8812 | -34.0579 | -5.159
2.8873 | -34.0582 | -5.1698
2.8934 | -34.0585 | -5.1806
2.8995 | -34.0588 | -5.1916
2.9056 | -34.0591 | -5.2026
2.9117 | -34.0594 | -5.2135
2.9178 | -34.0598 | -5.224
2.9239 | -34.0602 | -5.2354
2.93 | -34.0606 | -5.2467
2.9361 | -34.0609 | -5.2574
2.9423 | -34.0612 | -5.2679
2.9484 | -34.0616 | -5.2788
2.9545 | -34.062 | -5.2897
2.9606 | -34.0624 | -5.3009
2.9667 | -34.0627 | -5.3119
2.9728 | -34.063 | -5.3231
2.9789 | -34.0635 | -5.334
2.985 | -34.0639 | -5.3452
2.9911 | -34.0643 | -5.3564
2.9972 | -34.0647 | -5.3673
3.0033 | -34.0651 | -5.3787
3.0094 | -34.0655 | -5.3896
3.0155 | -34.0659 | -5.4004
3.0216 | -34.0663 | -5.4115
3.0277 | -34.0666 | -5.4231
3.0338 | -34.067 | -5.4343
3.0399 | -34.0674 | -5.4455
3.046 | -34.0677 | -5.4569
3.0521 | -34.068 | -5.4679
3.0582 | -34.0683 | -5.4799
3.0643 | -34.0686 | -5.4912
3.0704 | -34.069 | -5.5029
3.0765 | -34.0692 | -5.5139
3.0827 | -34.0695 | -5.525
3.0888 | -34.0699 | -5.5361
3.0949 | -34.0703 | -5.5474
3.101 | -34.0706 | -5.5587
3.1071 | -34.0709 | -5.5703
3.1132 | -34.0713 | -5.582
3.1193 | -34.0716 | -5.5933
3.1254 | -34.072 | -5.6046
3.1315 | -34.0724 | -5.6164
3.1376 | -34.0726 | -5.6277
3.1437 | -34.073 | -5.6391
3.1498 | -34.0733 | -5.651
3.1559 | -34.0736 | -5.6624
3.162 | -34.0739 | -5.6744
3.1681 | -34.0741 | -5.6863
3.1742 | -34.0745 | -5.6977
3.1803 | -34.0748 | -5.7093
3.1864 | -34.0751 | -5.7205
3.1925 | -34.0754 | -5.7319
3.1986 | -34.0757 | -5.7438
3.2047 | -34.076 | -5.7558
3.2108 | -34.0762 | -5.7668
3.2169 | -34.0765 | -5.7784
3.223 | -34.0768 | -5.7901
3.2292 | -34.077 | -5.8018
3.2353 | -34.0772 | -5.8125
3.2414 | -34.0774 | -5.8239
3.2475 | -34.0776 | -5.8352
3.2536 | -34.0779 | -5.8465
3.2597 | -34.0782 | -5.8574
3.2658 | -34.0784 | -5.8682
3.2719 | -34.0786 | -5.8792
3.278 | -34.0788 | -5.8907
3.2841 | -34.0791 | -5.9017
3.2902 | -34.0793 | -5.9131
3.2963 | -34.0796 | -5.9245
3.3024 | -34.0796 | -5.9355
3.3085 | -34.0799 | -5.9469
3.3146 | -34.0801 | -5.958
3.3207 | -34.0803 | -5.9694
3.3268 | -34.0805 | -5.9808
3.3329 | -34.0807 | -5.9914
3.339 | -34.0809 | -6.0022
3.3451 | -34.081 | -6.0133
3.3512 | -34.0812 | -6.0239
3.3573 | -34.0814 | -6.0351
3.3634 | -34.0816 | -6.0462
3.3696 | -34.0818 | -6.0574
3.3757 | -34.0819 | -6.0681
3.3818 | -34.0821 | -6.0789
3.3879 | -34.0823 | -6.0899
3.394 | -34.0824 | -6.1012
3.4001 | -34.0826 | -6.1119
3.4062 | -34.0827 | -6.1228
3.4123 | -34.0828 | -6.1333
3.4184 | -34.083 | -6.1445
3.4245 | -34.0832 | -6.1552
3.4306 | -34.0833 | -6.1658
3.4367 | -34.0835 | -6.1762
3.4428 | -34.0836 | -6.1865
3.4489 | -34.0837 | -6.1974
3.455 | -34.0839 | -6.2082
3.4611 | -34.0841 | -6.2193
3.4672 | -34.0842 | -6.2303
3.4733 | -34.0844 | -6.2413
3.4794 | -34.0845 | -6.252
3.4855 | -34.0846 | -6.2623
3.4916 | -34.0848 | -6.2726
3.4977 | -34.0849 | -6.2838
3.5038 | -34.0851 | -6.2946
3.5099 | -34.0852 | -6.3048
3.5161 | -34.0853 | -6.315
3.5222 | -34.0854 | -6.3253
3.5283 | -34.0855 | -6.336
3.5344 | -34.0856 | -6.3467
3.5405 | -34.0858 | -6.3571
3.5466 | -34.086 | -6.368
3.5527 | -34.086 | -6.3787
3.5588 | -34.0861 | -6.3892
3.5649 | -34.0862 | -6.3997
3.571 | -34.0863 | -6.4102
3.5771 | -34.0864 | -6.4213
3.5832 | -34.0865 | -6.431
3.5893 | -34.0868 | -6.4416
3.5954 | -34.0869 | -6.4523
3.6015 | -34.0871 | -6.4631
3.6076 | -34.0873 | -6.4735
3.6137 | -34.0874 | -6.4835
3.6198 | -34.0876 | -6.4937
3.6259 | -34.0878 | -6.5045
3.632 | -34.0879 | -6.515
3.6381 | -34.0881 | -6.5257
3.6442 | -34.0882 | -6.5362
3.6503 | -34.0884 | -6.5462
3.6565 | -34.0886 | -6.5569
3.6626 | -34.0887 | -6.5674
3.6687 | -34.0889 | -6.578
3.6748 | -34.0892 | -6.5883
3.6809 | -34.0894 | -6.5989
3.687 | -34.0896 | -6.6097
3.6931 | -34.0898 | -6.6199
3.6992 | -34.09 | -6.6304
3.7053 | -34.0902 | -6.6408
3.7114 | -34.0903 | -6.6514
3.7175 | -34.0905 | -6.662
3.7236 | -34.0908 | -6.6719
3.7297 | -34.091 | -6.6826
3.7358 | -34.0912 | -6.6934
3.7419 | -34.0913 | -6.7042
3.748 | -34.0916 | -6.7146
3.7541 | -34.0917 | -6.7251
3.7602 | -34.092 | -6.7355
3.7663 | -34.0922 | -6.7458
3.7724 | -34.0925 | -6.7564
3.7785 | -34.0927 | -6.7667
3.7846 | -34.0929 | -6.777
3.7907 | -34.0932 | -6.7879
3.7969 | -34.0935 | -6.7985
3.803 | -34.0937 | -6.8086
3.8091 | -34.0939 | -6.819
3.8152 | -34.0942 | -6.8294
3.8213 | -34.0944 | -6.8402
3.8274 | -34.0946 | -6.8504
3.8335 | -34.095 | -6.861
3.8396 | -34.0953 | -6.8715
3.8457 | -34.0956 | -6.8822
3.8518 | -34.0959 | -6.8926
3.8579 | -34.0962 | -6.903
3.864 | -34.0965 | -6.9137
3.8701 | -34.0968 | -6.9246
3.8762 | -34.097 | -6.935
3.8823 | -34.0973 | -6.9457
3.8884 | -34.0976 | -6.9563
3.8945 | -34.0979 | -6.9671
3.9006 | -34.0982 | -6.978
3.9067 | -34.0985 | -6.9889
3.9128 | -34.0988 | -6.9996
3.9189 | -34.0991 | -7.0099
3.925 | -34.0993 | -7.0206
3.9311 | -34.0996 | -7.0307
3.9372 | -34.0999 | -7.0415
3.9434 | -34.1002 | -7.0521
3.9495 | -34.1005 | -7.0628
3.9556 | -34.1009 | -7.0732
3.9617 | -34.1011 | -7.0834
3.9678 | -34.1015 | -7.0937
3.9739 | -34.1017 | -7.1043
3.98 | -34.102 | -7.115
3.9861 | -34.1023 | -7.1254
3.9922 | -34.1025 | -7.136
3.9983 | -34.1028 | -7.1462
4.0044 | -34.1031 | -7.1563
4.0105 | -34.1034 | -7.1669
4.0166 | -34.1037 | -7.1775
4.0227 | -34.104 | -7.188
4.0288 | -34.1043 | -7.1991
4.0349 | -34.1045 | -7.2094
4.041 | -34.1047 | -7.2199
4.0471 | -34.105 | -7.2305
4.0532 | -34.1053 | -7.2407
4.0593 | -34.1056 | -7.2514
4.0654 | -34.1059 | -7.2617
4.0715 | -34.1062 | -7.2727
4.0776 | -34.1065 | -7.2832
4.0838 | -34.1068 | -7.2941
4.0899 | -34.1071 | -7.3044
4.096 | -34.1073 | -7.3148
4.1021 | -34.1076 | -7.3251
4.1082 | -34.1078 | -7.3354
4.1143 | -34.1081 | -7.3456
4.1204 | -34.1084 | -7.3558
4.1265 | -34.1087 | -7.366
4.1326 | -34.1089 | -7.3761
4.1387 | -34.1093 | -7.3864
4.1448 | -34.1095 | -7.397
4.1509 | -34.1097 | -7.4075
4.157 | -34.11 | -7.4183
4.1631 | -34.1103 | -7.429
4.1692 | -34.1105 | -7.4393
4.1753 | -34.1108 | -7.4497
4.1814 | -34.111 | -7.4597
4.1875 | -34.1113 | -7.4705
4.1936 | -34.1116 | -7.481
4.1997 | -34.1118 | -7.4911
4.2058 | -34.112 | -7.5018
4.2119 | -34.1123 | -7.5121
4.218 | -34.1125 | -7.5222
4.2241 | -34.1127 | -7.5324
4.2303 | -34.113 | -7.5426
4.2364 | -34.1133 | -7.5529
4.2425 | -34.1134 | -7.5633
4.2486 | -34.1136 | -7.5739
4.2547 | -34.1138 | -7.584
4.2608 | -34.114 | -7.5945
4.2669 | -34.1142 | -7.6045
4.273 | -34.1144 | -7.615
4.2791 | -34.1147 | -7.6253
4.2852 | -34.1149 | -7.6358
4.2913 | -34.115 | -7.6459
4.2974 | -34.1153 | -7.6565
4.3035 | -34.1155 | -7.6671
4.3096 | -34.1157 | -7.6776
4.3157 | -34.1159 | -7.6877
4.3218 | -34.1161 | -7.698
4.3279 | -34.1164 | -7.7083
4.334 | -34.1165 | -7.7186
4.3401 | -34.1167 | -7.7284
4.3462 | -34.117 | -7.7388
4.3523 | -34.1173 | -7.7487
4.3584 | -34.1174 | -7.7586
4.3645 | -34.1176 | -7.7687
4.3707 | -34.1178 | -7.7787
4.3768 | -34.118 | -7.7889
4.3829 | -34.1183 | -7.7995
4.389 | -34.1184 | -7.8098
4.3951 | -34.1186 | -7.8198
4.4012 | -34.1188 | -7.8303
4.4073 | -34.119 | -7.8404
4.4134 | -34.1192 | -7.8505
4.4195 | -34.1195 | -7.8607
4.4256 | -34.1197 | -7.8707
4.4317 | -34.1199 | -7.8808
4.4378 | -34.1201 | -7.8907
4.4439 | -34.1203 | -7.9006
4.45 | -34.1204 | -7.9107
4.4561 | -34.1206 | -7.9209
4.4622 | -34.1207 | -7.9311
4.4683 | -34.121 | -7.9412
4.4744 | -34.1212 | -7.9517
4.4805 | -34.1214 | -7.9615
4.4866 | -34.1215 | -7.9716
4.4927 | -34.1218 | -7.982
4.4988 | -34.122 | -7.9921
4.5049 | -34.1222 | -8.0018
4.511 | -34.1224 | -8.0117
4.5172 | -34.1226 | -8.0217
4.5233 | -34.1228 | -8.0323
4.5294 | -34.123 | -8.0422
4.5355 | -34.1232 | -8.0524
4.5416 | -34.1234 | -8.0624
4.5477 | -34.1236 | -8.0725
4.5538 | -34.1237 | -8.0822
4.5599 | -34.124 | -8.0923
4.566 | -34.1242 | -8.1023
4.5721 | -34.1244 | -8.1125
4.5782 | -34.1246 | -8.1228
4.5843 | -34.1248 | -8.133
4.5904 | -34.1251 | -8.1431
4.5965 | -34.1253 | -8.1531
4.6026 | -34.1255 | -8.1637
4.6087 | -34.1257 | -8.1737
4.6148 | -34.1258 | -8.1841
4.6209 | -34.1261 | -8.1941
4.627 | -34.1263 | -8.2045
4.6331 | -34.1265 | -8.2144
4.6392 | -34.1267 | -8.2247
4.6453 | -34.1269 | -8.2347
4.6514 | -34.1271 | -8.2449
4.6576 | -34.1273 | -8.2554
4.6637 | -34.1274 | -8.2651
4.6698 | -34.1276 | -8.275
4.6759 | -34.1278 | -8.2855
4.682 | -34.128 | -8.2955
4.6881 | -34.1281 | -8.3056
4.6942 | -34.1284 | -8.3157
4.7003 | -34.1286 | -8.3258
4.7064 | -34.1288 | -8.3357
4.7125 | -34.129 | -8.3461
4.7186 | -34.1292 | -8.356
4.7247 | -34.1294 | -8.3662
4.7308 | -34.1297 | -8.3762
4.7369 | -34.1299 | -8.3864
4.743 | -34.1301 | -8.3969
4.7491 | -34.1303 | -8.4073
4.7552 | -34.1305 | -8.4176
4.7613 | -34.1306 | -8.4278
4.7674 | -34.1308 | -8.4378
4.7735 | -34.131 | -8.4484
4.7796 | -34.1313 | -8.4588
4.7857 | -34.1315 | -8.4691
4.7918 | -34.1317 | -8.4791
4.7979 | -34.132 | -8.4891
4.8041 | -34.1321 | -8.4993
4.8102 | -34.1324 | -8.5091
4.8163 | -34.1326 | -8.5192
4.8224 | -34.1328 | -8.5294
4.8285 | -34.133 | -8.5393
4.8346 | -34.1333 | -8.5492
4.8407 | -34.1334 | -8.5595
4.8468 | -34.1337 | -8.5699
4.8529 | -34.1339 | -8.5804
4.859 | -34.1341 | -8.5909
4.8651 | -34.1344 | -8.6008
4.8712 | -34.1346 | -8.6107
4.8773 | -34.1348 | -8.621
4.8834 | -34.135 | -8.6312
4.8895 | -34.1352 | -8.6411
4.8956 | -34.1354 | -8.6513
4.9017 | -34.1356 | -8.6612
4.9078 | -34.1358 | -8.6717
4.9139 | -34.136 | -8.6815
4.92 | -34.1363 | -8.6914
4.9261 | -34.1366 | -8.7012
4.9322 | -34.1368 | -8.7114
4.9383 | -34.137 | -8.7215
4.9445 | -34.1373 | -8.7317
4.9506 | -34.1375 | -8.7416
4.9567 | -34.1377 | -8.7516
4.9628 | -34.1379 | -8.7617
4.9689 | -34.1381 | -8.7722
4.975 | -34.1383 | -8.7826
4.9811 | -34.1386 | -8.7925
4.9872 | -34.1389 | -8.8025
4.9933 | -34.1392 | -8.813
4.9994 | -34.1394 | -8.8235
5.0055 | -34.1396 | -8.8338
5.0116 | -34.1399 | -8.8444
5.0177 | -34.1402 | -8.8548
5.0238 | -34.1404 | -8.8653
5.0299 | -34.1408 | -8.8759
5.036 | -34.141 | -8.8861
5.0421 | -34.1412 | -8.8962
5.0482 | -34.1415 | -8.9064
5.0543 | -34.1417 | -8.9164
5.0604 | -34.142 | -8.9266
5.0665 | -34.1423 | -8.9365
5.0726 | -34.1425 | -8.947
5.0787 | -34.1428 | -8.9575
5.0848 | -34.1431 | -8.9678
5.091 | -34.1433 | -8.9778
5.0971 | -34.1436 | -8.9877
5.1032 | -34.1438 | -8.9982
5.1093 | -34.1441 | -9.0084
5.1154 | -34.1443 | -9.0184
5.1215 | -34.1445 | -9.0284
5.1276 | -34.1448 | -9.0384
5.1337 | -34.1451 | -9.0483
5.1398 | -34.1454 | -9.0583
5.1459 | -34.1456 | -9.0688
5.152 | -34.1459 | -9.0792
5.1581 | -34.1461 | -9.0896
5.1642 | -34.1464 | -9.0999
5.1703 | -34.1467 | -9.1101
5.1764 | -34.147 | -9.1202
5.1825 | -34.1472 | -9.1302
5.1886 | -34.1475 | -9.1401
5.1947 | -34.1477 | -9.1503
5.2008 | -34.1481 | -9.1607
5.2069 | -34.1483 | -9.1712
5.213 | -34.1486 | -9.1812
5.2191 | -34.1489 | -9.1914
5.2252 | -34.1491 | -9.2017
5.2314 | -34.1494 | -9.2121
5.2375 | -34.1496 | -9.2221
5.2436 | -34.1499 | -9.2325
5.2497 | -34.1502 | -9.2424
5.2558 | -34.1505 | -9.253
5.2619 | -34.1507 | -9.2635
5.268 | -34.151 | -9.2739
5.2741 | -34.1513 | -9.2843
5.2802 | -34.1516 | -9.2949
5.2863 | -34.1519 | -9.3051
5.2924 | -34.1522 | -9.3156
5.2985 | -34.1525 | -9.3259
5.3046 | -34.1528 | -9.3363
5.3107 | -34.1531 | -9.3468
5.3168 | -34.1533 | -9.3571
5.3229 | -34.1536 | -9.3674
5.329 | -34.1539 | -9.3777
5.3351 | -34.1542 | -9.3877
5.3412 | -34.1544 | -9.398
5.3473 | -34.1547 | -9.4084
5.3534 | -34.155 | -9.4188
5.3595 | -34.1553 | -9.4292
5.3656 | -34.1556 | -9.4394
5.3717 | -34.1559 | -9.4501
5.3779 | -34.1562 | -9.4608
5.384 | -34.1565 | -9.4711
5.3901 | -34.1568 | -9.4817
5.3962 | -34.157 | -9.4918
5.4023 | -34.1573 | -9.5023
5.4084 | -34.1577 | -9.5126
5.4145 | -34.158 | -9.5227
5.4206 | -34.1582 | -9.5331
5.4267 | -34.1585 | -9.5438
5.4328 | -34.1588 | -9.5542
5.4389 | -34.1591 | -9.5643
5.445 | -34.1594 | -9.575
5.4511 | -34.1597 | -9.5851
5.4572 | -34.16 | -9.5955
5.4633 | -34.1603 | -9.6059
5.4694 | -34.1605 | -9.6163
5.4755 | -34.1608 | -9.6264
5.4816 | -34.1611 | -9.6367
5.4877 | -34.1614 | -9.647
5.4938 | -34.1617 | -9.6571
5.4999 | -34.162 | -9.6673
5.506 | -34.1624 | -9.6776
5.5121 | -34.1626 | -9.6879
5.5183 | -34.1629 | -9.6981
5.5244 | -34.1632 | -9.7086
5.5305 | -34.1635 | -9.7189
5.5366 | -34.1638 | -9.729
5.5427 | -34.164 | -9.7395
5.5488 | -34.1643 | -9.7499
5.5549 | -34.1646 | -9.7603
5.561 | -34.1649 | -9.7709
5.5671 | -34.1652 | -9.7811
5.5732 | -34.1655 | -9.7917
5.5793 | -34.1658 | -9.8024
5.5854 | -34.1661 | -9.8127
5.5915 | -34.1664 | -9.8234
5.5976 | -34.1667 | -9.8338
5.6037 | -34.167 | -9.8443
5.6098 | -34.1672 | -9.8546
5.6159 | -34.1676 | -9.8648
5.622 | -34.1678 | -9.8753
5.6281 | -34.1681 | -9.8857
5.6342 | -34.1684 | -9.8964
5.6403 | -34.1687 | -9.9068
5.6464 | -34.169 | -9.9174
5.6525 | -34.1693 | -9.9278
5.6586 | -34.1696 | -9.9379
5.6648 | -34.17 | -9.9483
5.6709 | -34.1703 | -9.9589
5.677 | -34.1706 | -9.9693
5.6831 | -34.1709 | -9.9798
5.6892 | -34.1711 | -9.9904
5.6953 | -34.1714 | -10.0008
5.7014 | -34.1718 | -10.0113
5.7075 | -34.172 | -10.0216
5.7136 | -34.1723 | -10.032
5.7197 | -34.1726 | -10.0424
5.7258 | -34.1729 | -10.0526
5.7319 | -34.1732 | -10.063
5.738 | -34.1735 | -10.0734
5.7441 | -34.1738 | -10.0837
5.7502 | -34.1742 | -10.0938
5.7563 | -34.1744 | -10.1042
5.7624 | -34.1748 | -10.1148
5.7685 | -34.1751 | -10.1253
5.7746 | -34.1753 | -10.1359
5.7807 | -34.1756 | -10.1464
5.7868 | -34.1759 | -10.157
5.7929 | -34.1761 | -10.1673
5.799 | -34.1764 | -10.1775
5.8052 | -34.1767 | -10.1878
5.8113 | -34.177 | -10.1981
5.8174 | -34.1773 | -10.2083
5.8235 | -34.1776 | -10.2187
5.8296 | -34.1779 | -10.2293
5.8357 | -34.1781 | -10.2396
5.8418 | -34.1784 | -10.2501
5.8479 | -34.1787 | -10.2606
5.854 | -34.179 | -10.2711
5.8601 | -34.1793 | -10.2813
5.8662 | -34.1796 | -10.2919
5.8723 | -34.1799 | -10.3021
5.8784 | -34.1801 | -10.3124
5.8845 | -34.1804 | -10.3226
5.8906 | -34.1807 | -10.3331
5.8967 | -34.181 | -10.3432
5.9028 | -34.1813 | -10.3534
5.9089 | -34.1817 | -10.3637
5.915 | -34.1819 | -10.3738
5.9211 | -34.1822 | -10.3841
5.9272 | -34.1825 | -10.3942
5.9333 | -34.1828 | -10.4045
5.9394 | -34.1831 | -10.4148
5.9455 | -34.1834 | -10.4249
5.9517 | -34.1838 | -10.4351
5.9578 | -34.1841 | -10.4455
5.9639 | -34.1844 | -10.4555
5.97 | -34.1847 | -10.4659
5.9761 | -34.1851 | -10.4762
5.9822 | -34.1853 | -10.4865
5.9883 | -34.1856 | -10.4968
5.9944 | -34.186 | -10.5073
6.0005 | -34.1863 | -10.5177
6.0066 | -34.1866 | -10.5278
6.0127 | -34.1869 | -10.5379
6.0188 | -34.1872 | -10.5485
6.0249 | -34.1875 | -10.5589
6.031 | -34.1877 | -10.569
6.0371 | -34.1881 | -10.5793
6.0432 | -34.1883 | -10.5895
6.0493 | -34.1887 | -10.5996
6.0554 | -34.189 | -10.6097
6.0615 | -34.1893 | -10.62
6.0676 | -34.1896 | -10.6305
6.0737 | -34.19 | -10.6409
6.0798 | -34.1902 | -10.6511
6.0859 | -34.1906 | -10.6613
6.0921 | -34.1909 | -10.6714
6.0982 | -34.1912 | -10.6816
6.1043 | -34.1915 | -10.6918
6.1104 | -34.1918 | -10.7022
6.1165 | -34.1921 | -10.712
6.1226 | -34.1925 | -10.7222
6.1287 | -34.1928 | -10.7324
6.1348 | -34.1931 | -10.7426
6.1409 | -34.1934 | -10.7528
6.147 | -34.1937 | -10.763
6.1531 | -34.194 | -10.7731
6.1592 | -34.1943 | -10.7832
6.1653 | -34.1946 | -10.7933
6.1714 | -34.1948 | -10.8035
6.1775 | -34.1952 | -10.8139
6.1836 | -34.1955 | -10.8241
6.1897 | -34.1958 | -10.8343
6.1958 | -34.1961 | -10.8448
6.2019 | -34.1964 | -10.8551
6.208 | -34.1968 | -10.8653
6.2141 | -34.1971 | -10.8755
6.2202 | -34.1974 | -10.8857
6.2263 | -34.1977 | -10.8958
6.2325 | -34.198 | -10.9057
6.2386 | -34.1983 | -10.9156
6.2447 | -34.1986 | -10.926
6.2508 | -34.199 | -10.9362
6.2569 | -34.1993 | -10.9463
6.263 | -34.1996 | -10.9565
6.2691 | -34.1999 | -10.9666
6.2752 | -34.2002 | -10.9768
6.2813 | -34.2005 | -10.987
6.2874 | -34.2008 | -10.997
6.2935 | -34.2011 | -11.0071
6.2996 | -34.2014 | -11.0172
6.3057 | -34.2017 | -11.0274
6.3118 | -34.2019 | -11.0377
6.3179 | -34.2022 | -11.0478
6.324 | -34.2026 | -11.0579
6.3301 | -34.2029 | -11.068
6.3362 | -34.2032 | -11.0781
6.3423 | -34.2035 | -11.0883
6.3484 | -34.2038 | -11.0986
6.3545 | -34.2041 | -11.1083
6.3606 | -34.2044 | -11.1182
6.3667 | -34.2047 | -11.1281
6.3728 | -34.205 | -11.1383
6.379 | -34.2053 | -11.1485
6.3851 | -34.2056 | -11.1583
6.3912 | -34.2059 | -11.1684
6.3973 | -34.2062 | -11.1785
6.4034 | -34.2065 | -11.1884
6.4095 | -34.2068 | -11.1985
6.4156 | -34.2071 | -11.2086
6.4217 | -34.2074 | -11.2186
6.4278 | -34.2077 | -11.2288
6.4339 | -34.208 | -11.2388
6.44 | -34.2083 | -11.2488
6.4461 | -34.2086 | -11.2591
6.4522 | -34.2089 | -11.269
6.4583 | -34.2092 | -11.2792
6.4644 | -34.2095 | -11.2894
6.4705 | -34.2098 | -11.2991
6.4766 | -34.2101 | -11.3092
6.4827 | -34.2104 | -11.3193
6.4888 | -34.2106 | -11.3292
6.4949 | -34.211 | -11.339
6.501 | -34.2113 | -11.3491
6.5071 | -34.2116 | -11.3589
6.5132 | -34.2119 | -11.369
6.5194 | -34.2122 | -11.3791
6.5255 | -34.2125 | -11.3889
6.5316 | -34.2127 | -11.399
6.5377 | -34.213 | -11.409
6.5438 | -34.2133 | -11.4188
6.5499 | -34.2136 | -11.4288
6.556 | -34.2139 | -11.4387
6.5621 | -34.2142 | -11.449
6.5682 | -34.2145 | -11.4591
6.5743 | -34.2149 | -11.469
6.5804 | -34.2151 | -11.4791
6.5865 | -34.2154 | -11.4891
6.5926 | -34.2158 | -11.4991
6.5987 | -34.2161 | -11.5088
6.6048 | -34.2164 | -11.5188
6.6109 | -34.2167 | -11.5288
6.617 | -34.2169 | -11.5389
6.6231 | -34.2172 | -11.5489
6.6292 | -34.2175 | -11.5588
6.6353 | -34.2178 | -11.5691
6.6414 | -34.2182 | -11.579
6.6475 | -34.2185 | -11.589
6.6536 | -34.2187 | -11.5988
6.6597 | -34.219 | -11.6085
6.6659 | -34.2193 | -11.6185
6.672 | -34.2197 | -11.6284
6.6781 | -34.2199 | -11.6384
6.6842 | -34.2203 | -11.6482
6.6903 | -34.2206 | -11.6581
6.6964 | -34.2209 | -11.6681
6.7025 | -34.2212 | -11.6779
6.7086 | -34.2215 | -11.6878
6.7147 | -34.2218 | -11.6978
6.7208 | -34.2221 | -11.7079
6.7269 | -34.2224 | -11.7178
6.733 | -34.2226 | -11.7275
6.7391 | -34.223 | -11.7374
6.7452 | -34.2233 | -11.7471
6.7513 | -34.2236 | -11.7567
6.7574 | -34.2239 | -11.7667
6.7635 | -34.2242 | -11.7765
6.7696 | -34.2245 | -11.7862
6.7757 | -34.2248 | -11.7962
6.7818 | -34.2251 | -11.8061
6.7879 | -34.2255 | -11.8158
6.794 | -34.2257 | -11.8255
6.8001 | -34.226 | -11.8353
6.8063 | -34.2263 | -11.8452
6.8124 | -34.2266 | -11.8552
6.8185 | -34.2269 | -11.8653
6.8246 | -34.2272 | -11.875
6.8307 | -34.2275 | -11.8847
6.8368 | -34.2278 | -11.8948
6.8429 | -34.2281 | -11.9047
6.849 | -34.2284 | -11.9147
6.8551 | -34.2288 | -11.9244
6.8612 | -34.229 | -11.9342
6.8673 | -34.2293 | -11.944
6.8734 | -34.2296 | -11.954
6.8795 | -34.2299 | -11.9639
6.8856 | -34.2302 | -11.9738
6.8917 | -34.2305 | -11.9836
6.8978 | -34.2308 | -11.9931
6.9039 | -34.2311 | -12.0027
6.91 | -34.2315 | -12.0126
6.9161 | -34.2318 | -12.0222
6.9222 | -34.2321 | -12.032
6.9283 | -34.2324 | -12.0416
6.9344 | -34.2327 | -12.0515
6.9405 | -34.2331 | -12.0613
6.9466 | -34.2334 | -12.0713
6.9528 | -34.2337 | -12.0813
6.9589 | -34.234 | -12.0912
6.965 | -34.2344 | -12.1011
6.9711 | -34.2347 | -12.1111
6.9772 | -34.235 | -12.1209
6.9833 | -34.2353 | -12.1307
6.9894 | -34.2356 | -12.1407
6.9955 | -34.236 | -12.1503
7.0016 | -34.2363 | -12.1601
7.0077 | -34.2366 | -12.1699
7.0138 | -34.2369 | -12.1796
7.0199 | -34.2372 | -12.1893
7.026 | -34.2375 | -12.1991
7.0321 | -34.2378 | -12.2086
7.0382 | -34.2382 | -12.2181
7.0443 | -34.2385 | -12.2281
7.0504 | -34.2388 | -12.238
7.0565 | -34.2391 | -12.2478
7.0626 | -34.2395 | -12.2573
7.0687 | -34.2398 | -12.2667
7.0748 | -34.2401 | -12.2762
7.0809 | -34.2404 | -12.286
7.087 | -34.2408 | -12.2958
7.0932 | -34.241 | -12.3056
7.0993 | -34.2414 | -12.3153
7.1054 | -34.2417 | -12.3249
7.1115 | -34.242 | -12.3344
7.1176 | -34.2423 | -12.3441
7.1237 | -34.2426 | -12.3538
7.1298 | -34.243 | -12.3634
7.1359 | -34.2433 | -12.3732
7.142 | -34.2436 | -12.3826
7.1481 | -34.2439 | -12.3923
7.1542 | -34.2442 | -12.402
7.1603 | -34.2445 | -12.4117
7.1664 | -34.2449 | -12.4214
7.1725 | -34.2452 | -12.4312
7.1786 | -34.2456 | -12.4409
7.1847 | -34.2459 | -12.4506
7.1908 | -34.2462 | -12.4605
7.1969 | -34.2465 | -12.4704
7.203 | -34.2468 | -12.4799
7.2091 | -34.2471 | -12.4896
7.2152 | -34.2475 | -12.4993
7.2213 | -34.2478 | -12.5088
7.2274 | -34.2481 | -12.5185
7.2335 | -34.2484 | -12.5281
7.2397 | -34.2488 | -12.5378
7.2458 | -34.2491 | -12.5472
7.2519 | -34.2495 | -12.5569
7.258 | -34.2498 | -12.5668
7.2641 | -34.2501 | -12.5764
7.2702 | -34.2504 | -12.5863
7.2763 | -34.2508 | -12.596
7.2824 | -34.2511 | -12.6059
7.2885 | -34.2514 | -12.6156
7.2946 | -34.2518 | -12.6253
7.3007 | -34.252 | -12.6351
7.3068 | -34.2524 | -12.6451
7.3129 | -34.2527 | -12.6549
7.319 | -34.253 | -12.6646
7.3251 | -34.2534 | -12.6741
7.3312 | -34.2537 | -12.6838
7.3373 | -34.254 | -12.6933
7.3434 | -34.2544 | -12.7029
7.3495 | -34.2547 | -12.7125
7.3556 | -34.2551 | -12.7222
7.3617 | -34.2554 | -12.7321
7.3678 | -34.2557 | -12.7417
7.3739 | -34.256 | -12.7516
7.3801 | -34.2564 | -12.7616
7.3862 | -34.2567 | -12.7712
7.3923 | -34.2571 | -12.781
7.3984 | -34.2574 | -12.7905
7.4045 | -34.2577 | -12.8001
7.4106 | -34.2581 | -12.8098
7.4167 | -34.2584 | -12.8196
7.4228 | -34.2588 | -12.8294
7.4289 | -34.2591 | -12.839
7.435 | -34.2594 | -12.8485
7.4411 | -34.2598 | -12.8583
7.4472 | -34.2601 | -12.8681
7.4533 | -34.2604 | -12.8776
7.4594 | -34.2608 | -12.8875
7.4655 | -34.2611 | -12.897
7.4716 | -34.2615 | -12.9069
7.4777 | -34.2618 | -12.9166
7.4838 | -34.2621 | -12.9261
7.4899 | -34.2624 | -12.9358
7.496 | -34.2628 | -12.9453
7.5021 | -34.2631 | -12.955
7.5082 | -34.2635 | -12.9646
7.5143 | -34.2638 | -12.9743
7.5204 | -34.2641 | -12.9842
7.5266 | -34.2645 | -12.9938
7.5327 | -34.2648 | -13.0034
7.5388 | -34.2651 | -13.0134
7.5449 | -34.2655 | -13.0232
7.551 | -34.2658 | -13.0328
7.5571 | -34.2662 | -13.0426
7.5632 | -34.2665 | -13.0524
7.5693 | -34.2669 | -13.062
7.5754 | -34.2672 | -13.0715
7.5815 | -34.2676 | -13.0812
7.5876 | -34.2679 | -13.0912
7.5937 | -34.2683 | -13.1011
7.5998 | -34.2687 | -13.1107
7.6059 | -34.269 | -13.1206
7.612 | -34.2693 | -13.1302
7.6181 | -34.2697 | -13.1399
7.6242 | -34.27 | -13.1499
7.6303 | -34.2704 | -13.1598
7.6364 | -34.2707 | -13.1696
7.6425 | -34.271 | -13.1793
7.6486 | -34.2713 | -13.1891
7.6547 | -34.2717 | -13.1987
7.6608 | -34.2721 | -13.2084
7.667 | -34.2724 | -13.2182
7.6731 | -34.2727 | -13.228
7.6792 | -34.2731 | -13.238
7.6853 | -34.2734 | -13.2478
7.6914 | -34.2738 | -13.2577
7.6975 | -34.2741 | -13.2675
7.7036 | -34.2745 | -13.2773
7.7097 | -34.2748 | -13.2872
7.7158 | -34.2752 | -13.2971
7.7219 | -34.2756 | -13.3067
7.728 | -34.2759 | -13.3166
7.7341 | -34.2762 | -13.3265
7.7402 | -34.2766 | -13.3363
7.7463 | -34.2769 | -13.3461
7.7524 | -34.2773 | -13.3561
7.7585 | -34.2776 | -13.3659
7.7646 | -34.278 | -13.3757
7.7707 | -34.2783 | -13.3857
7.7768 | -34.2786 | -13.3956
7.7829 | -34.279 | -13.4057
7.789 | -34.2793 | -13.4157
7.7951 | -34.2797 | -13.4255
7.8012 | -34.28 | -13.4353
7.8073 | -34.2804 | -13.4453
7.8135 | -34.2807 | -13.4551
7.8196 | -34.2811 | -13.4652
7.8257 | -34.2815 | -13.4752
7.8318 | -34.2818 | -13.4851
7.8379 | -34.2821 | -13.4948
7.844 | -34.2825 | -13.5046
7.8501 | -34.2828 | -13.5145
7.8562 | -34.2832 | -13.5246
7.8623 | -34.2836 | -13.5343
7.8684 | -34.2839 | -13.5442
7.8745 | -34.2842 | -13.5541
7.8806 | -34.2846 | -13.5642
7.8867 | -34.2849 | -13.5738
7.8928 | -34.2853 | -13.5834
7.8989 | -34.2856 | -13.5932
7.905 | -34.286 | -13.6032
7.9111 | -34.2863 | -13.6131
7.9172 | -34.2867 | -13.6229
7.9233 | -34.287 | -13.6328
7.9294 | -34.2874 | -13.6428
7.9355 | -34.2877 | -13.6524
7.9416 | -34.288 | -13.6623
7.9477 | -34.2884 | -13.672
7.9539 | -34.2887 | -13.6819
7.96 | -34.2891 | -13.6918
7.9661 | -34.2894 | -13.7018
7.9722 | -34.2898 | -13.7115
7.9783 | -34.2902 | -13.7216
7.9844 | -34.2906 | -13.7314
7.9905 | -34.2909 | -13.7415
7.9966 | -34.2912 | -13.7515
8.0027 | -34.2916 | -13.7613
8.0088 | -34.2919 | -13.7712
8.0149 | -34.2922 | -13.7811
8.021 | -34.2926 | -13.791
8.0271 | -34.2929 | -13.8012
8.0332 | -34.2933 | -13.811
8.0393 | -34.2937 | -13.821
8.0454 | -34.294 | -13.8312
8.0515 | -34.2944 | -13.841
8.0576 | -34.2947 | -13.8512
8.0637 | -34.2951 | -13.861
8.0698 | -34.2954 | -13.8707
8.0759 | -34.2957 | -13.8805
8.082 | -34.2961 | -13.8903
8.0881 | -34.2964 | -13.9003
8.0942 | -34.2968 | -13.9102
8.1004 | -34.2972 | -13.9202
8.1065 | -34.2975 | -13.93
8.1126 | -34.2978 | -13.9399
8.1187 | -34.2982 | -13.9499
8.1248 | -34.2985 | -13.9596
8.1309 | -34.2989 | -13.9693
8.137 | -34.2992 | -13.979
8.1431 | -34.2996 | -13.9892
8.1492 | -34.2999 | -13.999
8.1553 | -34.3003 | -14.009
8.1614 | -34.3006 | -14.019
8.1675 | -34.301 | -14.0289
8.1736 | -34.3013 | -14.039
8.1797 | -34.3016 | -14.0489
8.1858 | -34.302 | -14.0589
8.1919 | -34.3023 | -14.0688
8.198 | -34.3027 | -14.0785
8.2041 | -34.3031 | -14.0885
8.2102 | -34.3034 | -14.0983
8.2163 | -34.3038 | -14.1081
8.2224 | -34.3041 | -14.1179
8.2285 | -34.3045 | -14.1276
8.2346 | -34.3048 | -14.1374
8.2408 | -34.3052 | -14.1472
8.2469 | -34.3055 | -14.1571
8.253 | -34.3059 | -14.1669
8.2591 | -34.3062 | -14.1768
8.2652 | -34.3066 | -14.1868
8.2713 | -34.307 | -14.1967
8.2774 | -34.3073 | -14.2067
8.2835 | -34.3077 | -14.2168
8.2896 | -34.308 | -14.2266
8.2957 | -34.3084 | -14.2365
8.3018 | -34.3088 | -14.2464
8.3079 | -34.3091 | -14.2561
8.314 | -34.3095 | -14.266
8.3201 | -34.3099 | -14.2758
8.3262 | -34.3102 | -14.2857
8.3323 | -34.3105 | -14.2956
8.3384 | -34.3109 | -14.3057
8.3445 | -34.3112 | -14.3154
8.3506 | -34.3116 | -14.3254
8.3567 | -34.312 | -14.3354
8.3628 | -34.3123 | -14.3452
8.3689 | -34.3127 | -14.3549
8.375 | -34.313 | -14.3649
8.3812 | -34.3133 | -14.3747
8.3873 | -34.3137 | -14.3846
8.3934 | -34.3141 | -14.3946
8.3995 | -34.3144 | -14.4043
8.4056 | -34.3148 | -14.4142
8.4117 | -34.3151 | -14.4241
8.4178 | -34.3154 | -14.434
8.4239 | -34.3158 | -14.4437
8.43 | -34.3161 | -14.4536
8.4361 | -34.3164 | -14.4636
8.4422 | -34.3168 | -14.4733
8.4483 | -34.3172 | -14.4831
8.4544 | -34.3175 | -14.493
8.4605 | -34.3179 | -14.5027
8.4666 | -34.3182 | -14.5128
8.4727 | -34.3186 | -14.5227
8.4788 | -34.3189 | -14.5325
8.4849 | -34.3193 | -14.5424
8.491 | -34.3196 | -14.5522
8.4971 | -34.32 | -14.5618
8.5032 | -34.3203 | -14.5716
8.5093 | -34.3207 | -14.5812
8.5154 | -34.321 | -14.5912
8.5215 | -34.3214 | -14.6009
8.5277 | -34.3217 | -14.6107
8.5338 | -34.3221 | -14.6205
8.5399 | -34.3224 | -14.6303
8.546 | -34.3228 | -14.6399
8.5521 | -34.3232 | -14.6498
8.5582 | -34.3235 | -14.6595
8.5643 | -34.3238 | -14.6691
8.5704 | -34.3242 | -14.6787
8.5765 | -34.3245 | -14.6886
8.5826 | -34.3249 | -14.6981
8.5887 | -34.3253 | -14.7078
8.5948 | -34.3256 | -14.7176
8.6009 | -34.326 | -14.7272
8.607 | -34.3263 | -14.7372
8.6131 | -34.3267 | -14.747
8.6192 | -34.3271 | -14.7566
8.6253 | -34.3274 | -14.7663
8.6314 | -34.3278 | -14.776
8.6375 | -34.3281 | -14.7858
8.6436 | -34.3285 | -14.7956
8.6497 | -34.3288 | -14.8052
8.6558 | -34.3291 | -14.815
8.6619 | -34.3295 | -14.8247
8.6681 | -34.3298 | -14.8346
8.6742 | -34.3302 | -14.8444
8.6803 | -34.3305 | -14.8541
8.6864 | -34.3309 | -14.8636
8.6925 | -34.3312 | -14.8732
8.6986 | -34.3316 | -14.883
8.7047 | -34.3319 | -14.8929
8.7108 | -34.3323 | -14.9026
8.7169 | -34.3326 | -14.9124
8.723 | -34.333 | -14.9219
8.7291 | -34.3333 | -14.9315
8.7352 | -34.3337 | -14.9414
8.7413 | -34.334 | -14.9512
8.7474 | -34.3344 | -14.9608
8.7535 | -34.3347 | -14.9705
8.7596 | -34.3351 | -14.9802
8.7657 | -34.3355 | -14.99
8.7718 | -34.3358 | -14.9998
8.7779 | -34.3362 | -15.0097
8.784 | -34.3365 | -15.0194
8.7901 | -34.3369 | -15.0291
8.7962 | -34.3373 | -15.0386
8.8023 | -34.3376 | -15.0482
8.8084 | -34.3379 | -15.0579
8.8146 | -34.3383 | -15.0675
8.8207 | -34.3386 | -15.0772
8.8268 | -34.339 | -15.0865
8.8329 | -34.3394 | -15.0962
8.839 | -34.3397 | -15.1059
8.8451 | -34.3401 | -15.1155
8.8512 | -34.3405 | -15.1251
8.8573 | -34.3408 | -15.1348
8.8634 | -34.3411 | -15.1442
8.8695 | -34.3415 | -15.1541
8.8756 | -34.3418 | -15.1638
8.8817 | -34.3422 | -15.1733
8.8878 | -34.3425 | -15.1832
8.8939 | -34.3429 | -15.1928
8.9 | -34.3433 | -15.2024
8.9061 | -34.3436 | -15.2121
8.9122 | -34.344 | -15.2217
8.9183 | -34.3443 | -15.2314
8.9244 | -34.3447 | -15.2412
8.9305 | -34.3451 | -15.2506
8.9366 | -34.3455 | -15.2602
8.9427 | -34.3458 | -15.27
8.9488 | -34.3462 | -15.2795
8.955 | -34.3465 | -15.2891
8.9611 | -34.3469 | -15.2989
8.9672 | -34.3473 | -15.3084
8.9733 | -34.3477 | -15.3179
8.9794 | -34.3481 | -15.3274
8.9855 | -34.3484 | -15.3371
8.9916 | -34.3487 | -15.3469
8.9977 | -34.3491 | -15.3565
9.0038 | -34.3495 | -15.366
9.0099 | -34.3499 | -15.3756
9.016 | -34.3502 | -15.3853
9.0221 | -34.3506 | -15.3948
9.0282 | -34.351 | -15.4042
9.0343 | -34.3514 | -15.4139
9.0404 | -34.3518 | -15.4235
9.0465 | -34.3521 | -15.4333
9.0526 | -34.3525 | -15.4428
9.0587 | -34.3529 | -15.4524
9.0648 | -34.3533 | -15.4619
9.0709 | -34.3537 | -15.4714
9.077 | -34.3541 | -15.4811
9.0831 | -34.3544 | -15.4908
9.0892 | -34.3548 | -15.5005
9.0953 | -34.3552 | -15.5101
9.1015 | -34.3556 | -15.5197
9.1076 | -34.356 | -15.5291
9.1137 | -34.3563 | -15.5386
9.1198 | -34.3567 | -15.5482
9.1259 | -34.3571 | -15.5578
9.132 | -34.3575 | -15.5675
9.1381 | -34.3579 | -15.577
9.1442 | -34.3583 | -15.5863
9.1503 | -34.3586 | -15.596
9.1564 | -34.359 | -15.6055
9.1625 | -34.3594 | -15.6151
9.1686 | -34.3598 | -15.6246
9.1747 | -34.3602 | -15.6341
9.1808 | -34.3606 | -15.6436
9.1869 | -34.3609 | -15.6532
9.193 | -34.3613 | -15.6628
9.1991 | -34.3617 | -15.6724
9.2052 | -34.3621 | -15.682
9.2113 | -34.3625 | -15.6915
9.2174 | -34.3629 | -15.701
9.2235 | -34.3633 | -15.7105
9.2296 | -34.3637 | -15.72
9.2357 | -34.364 | -15.7298
9.2419 | -34.3644 | -15.7393
9.248 | -34.3649 | -15.7488
9.2541 | -34.3653 | -15.7584
9.2602 | -34.3657 | -15.7681
9.2663 | -34.366 | -15.7774
9.2724 | -34.3664 | -15.7869
9.2785 | -34.3668 | -15.7964
9.2846 | -34.3672 | -15.8058
9.2907 | -34.3676 | -15.8153
9.2968 | -34.368 | -15.825
9.3029 | -34.3684 | -15.8347
9.309 | -34.3688 | -15.8442
9.3151 | -34.3691 | -15.8537
9.3212 | -34.3696 | -15.863
9.3273 | -34.37 | -15.8725
9.3334 | -34.3704 | -15.8819
9.3395 | -34.3707 | -15.8913
9.3456 | -34.3711 | -15.9007
9.3517 | -34.3715 | -15.9103
9.3578 | -34.3719 | -15.9198
9.3639 | -34.3723 | -15.9291
9.37 | -34.3727 | -15.9387
9.3761 | -34.3731 | -15.9484
9.3822 | -34.3735 | -15.9579
9.3884 | -34.3739 | -15.9671
9.3945 | -34.3743 | -15.9767
9.4006 | -34.3747 | -15.9862
9.4067 | -34.3751 | -15.9958
9.4128 | -34.3755 | -16.0054
9.4189 | -34.3759 | -16.0151
9.425 | -34.3763 | -16.0244
9.4311 | -34.3767 | -16.034
9.4372 | -34.3772 | -16.0434
9.4433 | -34.3776 | -16.0529
9.4494 | -34.3779 | -16.0622
9.4555 | -34.3784 | -16.0718
9.4616 | -34.3788 | -16.0813
9.4677 | -34.3792 | -16.0908
9.4738 | -34.3796 | -16.1003
9.4799 | -34.38 | -16.1098
9.486 | -34.3804 | -16.1193
9.4921 | -34.3808 | -16.1288
9.4982 | -34.3812 | -16.1381
9.5043 | -34.3816 | -16.1476
9.5104 | -34.382 | -16.1569
9.5165 | -34.3824 | -16.1663
9.5226 | -34.3828 | -16.176
9.5288 | -34.3832 | -16.1853
9.5349 | -34.3836 | -16.1948
9.541 | -34.384 | -16.2041
9.5471 | -34.3844 | -16.2136
9.5532 | -34.3848 | -16.2231
9.5593 | -34.3852 | -16.2327
9.5654 | -34.3856 | -16.2421
9.5715 | -34.386 | -16.2515
9.5776 | -34.3864 | -16.2611
9.5837 | -34.3868 | -16.2706
9.5898 | -34.3871 | -16.2802
9.5959 | -34.3876 | -16.2897
9.602 | -34.388 | -16.299
9.6081 | -34.3884 | -16.3084
9.6142 | -34.3888 | -16.3177
9.6203 | -34.3892 | -16.3272
9.6264 | -34.3896 | -16.3366
9.6325 | -34.3899 | -16.3461
9.6386 | -34.3903 | -16.3554
9.6447 | -34.3907 | -16.3649
9.6508 | -34.3911 | -16.3743
9.6569 | -34.3915 | -16.3837
9.663 | -34.392 | -16.3932
9.6691 | -34.3924 | -16.4027
9.6753 | -34.3928 | -16.4121
9.6814 | -34.3932 | -16.4214
9.6875 | -34.3936 | -16.4307
9.6936 | -34.394 | -16.44
9.6997 | -34.3944 | -16.4495
9.7058 | -34.3949 | -16.4589
9.7119 | -34.3953 | -16.4685
9.718 | -34.3957 | -16.4781
9.7241 | -34.3961 | -16.4874
9.7302 | -34.3964 | -16.4969
9.7363 | -34.3969 | -16.5063
9.7424 | -34.3973 | -16.5155
9.7485 | -34.3977 | -16.5251
9.7546 | -34.3981 | -16.5343
9.7607 | -34.3985 | -16.544
9.7668 | -34.3989 | -16.5535
9.7729 | -34.3993 | -16.5628
9.779 | -34.3997 | -16.5721
9.7851 | -34.4001 | -16.5817
9.7912 | -34.4006 | -16.591
9.7973 | -34.401 | -16.6006
9.8034 | -34.4014 | -16.6098
9.8095 | -34.4018 | -16.6194
9.8157 | -34.4023 | -16.6289
9.8218 | -34.4027 | -16.6385
9.8279 | -34.4031 | -16.6478
9.834 | -34.4034 | -16.6573
9.8401 | -34.4039 | -16.6667
9.8462 | -34.4043 | -16.676
9.8523 | -34.4047 | -16.6854
9.8584 | -34.4051 | -16.6948
9.8645 | -34.4055 | -16.7044
9.8706 | -34.4059 | -16.7138
9.8767 | -34.4063 | -16.7231
9.8828 | -34.4066 | -16.7325
9.8889 | -34.4071 | -16.7417
9.895 | -34.4075 | -16.7511
9.9011 | -34.4079 | -16.7605
9.9072 | -34.4083 | -16.7698
9.9133 | -34.4087 | -16.7792
9.9194 | -34.4091 | -16.7886
9.9255 | -34.4095 | -16.7979
9.9316 | -34.4099 | -16.8073
9.9377 | -34.4103 | -16.8167
9.9438 | -34.4107 | -16.8262
9.9499 | -34.4111 | -16.8358
9.956 | -34.4115 | -16.8451
9.9622 | -34.4119 | -16.8546
9.9683 | -34.4123 | -16.864
9.9744 | -34.4127 | -16.8735
9.9805 | -34.4131 | -16.8829
9.9866 | -34.4135 | -16.8922
9.9927 | -34.4139 | -16.9014
9.9988 | -34.4143 | -16.9109
10.0049 | -34.4148 | -16.9203
10.011 | -34.4152 | -16.9297
10.0171 | -34.4156 | -16.939
10.0232 | -34.416 | -16.9484
10.0293 | -34.4164 | -16.9578
10.0354 | -34.4168 | -16.9671
10.0415 | -34.4173 | -16.9766
10.0476 | -34.4176 | -16.9861
10.0537 | -34.418 | -16.9956
10.0598 | -34.4185 | -17.005
10.0659 | -34.4188 | -17.0144
10.072 | -34.4192 | -17.0235
10.0781 | -34.4196 | -17.0329
10.0842 | -34.4201 | -17.0422
10.0903 | -34.4205 | -17.0516
10.0964 | -34.4209 | -17.0608
10.1026 | -34.4213 | -17.0702
10.1087 | -34.4217 | -17.0796
10.1148 | -34.422 | -17.0891
10.1209 | -34.4224 | -17.0982
10.127 | -34.4228 | -17.1075
10.1331 | -34.4233 | -17.1168
10.1392 | -34.4237 | -17.1259
10.1453 | -34.4241 | -17.1352
10.1514 | -34.4245 | -17.1446
10.1575 | -34.4249 | -17.1541
10.1636 | -34.4253 | -17.1632
10.1697 | -34.4257 | -17.1727
10.1758 | -34.4261 | -17.1818
10.1819 | -34.4265 | -17.1912
10.188 | -34.4269 | -17.2005
10.1941 | -34.4273 | -17.2099
10.2002 | -34.4276 | -17.2193
10.2063 | -34.428 | -17.2286
10.2124 | -34.4284 | -17.238
10.2185 | -34.4289 | -17.2474
10.2246 | -34.4293 | -17.2567
10.2307 | -34.4297 | -17.2661
10.2368 | -34.4301 | -17.2752
10.2429 | -34.4305 | -17.2845
10.2491 | -34.4308 | -17.2938
10.2552 | -34.4312 | -17.3032
10.2613 | -34.4316 | -17.3126
10.2674 | -34.432 | -17.322
10.2735 | -34.4324 | -17.3315
10.2796 | -34.4329 | -17.3407
10.2857 | -34.4333 | -17.3503
10.2918 | -34.4337 | -17.3597
10.2979 | -34.4341 | -17.369
10.304 | -34.4345 | -17.3782
10.3101 | -34.4348 | -17.3876
10.3162 | -34.4353 | -17.397
10.3223 | -34.4356 | -17.4062
10.3284 | -34.436 | -17.4155
10.3345 | -34.4364 | -17.4249
10.3406 | -34.4368 | -17.434
10.3467 | -34.4372 | -17.4433
10.3528 | -34.4376 | -17.4524
10.3589 | -34.438 | -17.4617
10.365 | -34.4384 | -17.4711
10.3711 | -34.4388 | -17.4806
10.3772 | -34.4392 | -17.4897
10.3833 | -34.4396 | -17.4989
10.3895 | -34.44 | -17.5081
10.3956 | -34.4404 | -17.5175
10.4017 | -34.4408 | -17.5269
10.4078 | -34.4412 | -17.536
10.4139 | -34.4415 | -17.5452
10.42 | -34.442 | -17.5549
10.4261 | -34.4423 | -17.5641
10.4322 | -34.4427 | -17.5732
10.4383 | -34.4431 | -17.5824
10.4444 | -34.4435 | -17.5916
10.4505 | -34.4438 | -17.6008
10.4566 | -34.4442 | -17.6099
10.4627 | -34.4446 | -17.619
10.4688 | -34.445 | -17.6281
10.4749 | -34.4454 | -17.6375
10.481 | -34.4458 | -17.6467
10.4871 | -34.4462 | -17.656
10.4932 | -34.4466 | -17.6653
10.4993 | -34.4471 | -17.6745
10.5054 | -34.4474 | -17.6838
10.5115 | -34.4478 | -17.6931
10.5176 | -34.4482 | -17.7022
10.5237 | -34.4486 | -17.7117
10.5298 | -34.449 | -17.7209
10.536 | -34.4494 | -17.7304
10.5421 | -34.4498 | -17.7396
10.5482 | -34.4502 | -17.749
10.5543 | -34.4506 | -17.7583
10.5604 | -34.451 | -17.7674
10.5665 | -34.4514 | -17.7767
10.5726 | -34.4518 | -17.786
10.5787 | -34.4522 | -17.7953
10.5848 | -34.4525 | -17.8042
10.5909 | -34.4529 | -17.8135
10.597 | -34.4533 | -17.8227
10.6031 | -34.4537 | -17.8321
10.6092 | -34.454 | -17.8414
10.6153 | -34.4544 | -17.8506
10.6214 | -34.4548 | -17.8596
10.6275 | -34.4552 | -17.869
10.6336 | -34.4556 | -17.8782
10.6397 | -34.4559 | -17.8874
10.6458 | -34.4563 | -17.8965
10.6519 | -34.4567 | -17.9059
10.658 | -34.4571 | -17.9157
10.6641 | -34.4575 | -17.9254
10.6702 | -34.4579 | -17.9344
10.6764 | -34.4583 | -17.9437
10.6825 | -34.4587 | -17.953
10.6886 | -34.459 | -17.9621
10.6947 | -34.4594 | -17.9713
10.7008 | -34.4598 | -17.9806
10.7069 | -34.4601 | -17.99
10.713 | -34.4604 | -17.9994
10.7191 | -34.4608 | -18.0088
10.7252 | -34.4612 | -18.0179
10.7313 | -34.4616 | -18.0272
10.7374 | -34.462 | -18.0366
10.7435 | -34.4623 | -18.0459
10.7496 | -34.4627 | -18.055
10.7557 | -34.4631 | -18.0644
10.7618 | -34.4635 | -18.0732
10.7679 | -34.4639 | -18.0828
10.774 | -34.4642 | -18.0917
10.7801 | -34.4647 | -18.1012
10.7862 | -34.4651 | -18.1107
10.7923 | -34.4655 | -18.1205
10.7984 | -34.466 | -18.1294
10.8045 | -34.4663 | -18.1384
10.8106 | -34.4667 | -18.1478
10.8168 | -34.4671 | -18.1566
10.8229 | -34.4675 | -18.1658
10.829 | -34.4679 | -18.1743
10.8351 | -34.4683 | -18.1841
10.8412 | -34.4688 | -18.1929
10.8473 | -34.4691 | -18.2024
10.8534 | -34.4696 | -18.2118
10.8595 | -34.47 | -18.2212
10.8656 | -34.4704 | -18.2304
10.8717 | -34.4708 | -18.2395
10.8778 | -34.4712 | -18.249
10.8839 | -34.4715 | -18.258
10.89 | -34.4719 | -18.2675
10.8961 | -34.4724 | -18.2768
10.9022 | -34.4728 | -18.2859
10.9083 | -34.4731 | -18.2951
10.9144 | -34.4735 | -18.3044
10.9205 | -34.4739 | -18.3138
10.9266 | -34.4743 | -18.3224
10.9327 | -34.4747 | -18.3318
10.9388 | -34.4752 | -18.3413
10.9449 | -34.4756 | -18.3502
10.951 | -34.4759 | -18.3591
10.9571 | -34.4763 | -18.3678
10.9633 | -34.4767 | -18.3771
10.9694 | -34.4771 | -18.3859
10.9755 | -34.4775 | -18.3951
10.9816 | -34.4778 | -18.4044
10.9877 | -34.4781 | -18.4132
10.9938 | -34.4784 | -18.4225
10.9999 | -34.4788 | -18.431
11.006 | -34.4792 | -18.4396
11.0121 | -34.4796 | -18.4491
11.0182 | -34.4799 | -18.4585
11.0243 | -34.4803 | -18.4676
11.0304 | -34.4807 | -18.477
11.0365 | -34.4811 | -18.4857
11.0426 | -34.4816 | -18.4942
11.0487 | -34.482 | -18.5032
11.0548 | -34.4824 | -18.5118
11.0609 | -34.4828 | -18.5209
11.067 | -34.4831 | -18.5301
11.0731 | -34.4834 | -18.5392
11.0792 | -34.4837 | -18.5483
11.0853 | -34.4842 | -18.5577
11.0914 | -34.4846 | -18.5667
11.0975 | -34.485 | -18.5759
11.1037 | -34.4853 | -18.5852
11.1098 | -34.4857 | -18.5944
11.1159 | -34.4861 | -18.6027
11.122 | -34.4864 | -18.6117
11.1281 | -34.4868 | -18.621
11.1342 | -34.4872 | -18.63
11.1403 | -34.4876 | -18.6392
11.1464 | -34.488 | -18.6483
11.1525 | -34.4883 | -18.6571
11.1586 | -34.4887 | -18.666
11.1647 | -34.4891 | -18.6752
11.1708 | -34.4894 | -18.6844
11.1769 | -34.4898 | -18.694
11.183 | -34.4902 | -18.7022
11.1891 | -34.4906 | -18.7114
11.1952 | -34.4911 | -18.7203
11.2013 | -34.4915 | -18.7293
11.2074 | -34.4918 | -18.7385
11.2135 | -34.4922 | -18.7474
11.2196 | -34.4925 | -18.7563
11.2257 | -34.4929 | -18.7658
11.2318 | -34.4933 | -18.7741
11.2379 | -34.4936 | -18.7832
11.244 | -34.4941 | -18.7926
11.2502 | -34.4945 | -18.8015
11.2563 | -34.4949 | -18.8107
11.2624 | -34.4953 | -18.8194
11.2685 | -34.4956 | -18.8286
11.2746 | -34.496 | -18.8371
11.2807 | -34.4964 | -18.8461
11.2868 | -34.4968 | -18.8552
11.2929 | -34.4972 | -18.8636
11.299 | -34.4976 | -18.8729
11.3051 | -34.498 | -18.8821
11.3112 | -34.4984 | -18.891
11.3173 | -34.4988 | -18.8996
11.3234 | -34.4992 | -18.9082
11.3295 | -34.4995 | -18.9172
11.3356 | -34.4999 | -18.9262
11.3417 | -34.5003 | -18.9351
11.3478 | -34.5006 | -18.9441
11.3539 | -34.5009 | -18.9535
11.36 | -34.5014 | -18.962
11.3661 | -34.5018 | -18.9709
11.3722 | -34.5022 | -18.9798
11.3783 | -34.5026 | -18.9887
11.3844 | -34.503 | -18.9975
11.3906 | -34.5035 | -19.0065
11.3967 | -34.5039 | -19.0153
11.4028 | -34.5043 | -19.0245
11.4089 | -34.5046 | -19.0335
11.415 | -34.505 | -19.0427
11.4211 | -34.5054 | -19.0515
11.4272 | -34.5058 | -19.0604
11.4333 | -34.5062 | -19.0686
11.4394 | -34.5067 | -19.0775
11.4455 | -34.5071 | -19.086
11.4516 | -34.5075 | -19.0947
11.4577 | -34.5078 | -19.1036
11.4638 | -34.5082 | -19.1124
11.4699 | -34.5086 | -19.1215
11.476 | -34.509 | -19.1305
11.4821 | -34.5094 | -19.1393
11.4882 | -34.5098 | -19.1483
11.4943 | -34.5102 | -19.1573
11.5004 | -34.5106 | -19.166
11.5065 | -34.511 | -19.1747
11.5126 | -34.5114 | -19.1835
11.5187 | -34.5118 | -19.1924
11.5248 | -34.5122 | -19.2015
11.5309 | -34.5126 | -19.2103
11.5371 | -34.513 | -19.2188
11.5432 | -34.5133 | -19.2278
11.5493 | -34.5137 | -19.2366
11.5554 | -34.5141 | -19.2457
11.5615 | -34.5145 | -19.2548
11.5676 | -34.5149 | -19.2636
11.5737 | -34.5153 | -19.2722
11.5798 | -34.5156 | -19.2809
11.5859 | -34.5161 | -19.2897
11.592 | -34.5164 | -19.2987
11.5981 | -34.5168 | -19.3072
11.6042 | -34.5173 | -19.3159
11.6103 | -34.5176 | -19.3246
11.6164 | -34.518 | -19.3334
11.6225 | -34.5184 | -19.3421
11.6286 | -34.5188 | -19.3509
11.6347 | -34.5191 | -19.3598
11.6408 | -34.5196 | -19.3682
11.6469 | -34.52 | -19.377
11.653 | -34.5203 | -19.3858
11.6591 | -34.5207 | -19.3947
11.6652 | -34.5211 | -19.4032
11.6713 | -34.5215 | -19.412
11.6775 | -34.522 | -19.4205
11.6836 | -34.5224 | -19.4295
11.6897 | -34.5228 | -19.4383
11.6958 | -34.5232 | -19.4473
11.7019 | -34.5236 | -19.4563
11.708 | -34.524 | -19.4651
11.7141 | -34.5244 | -19.4737
11.7202 | -34.5248 | -19.4824
11.7263 | -34.5252 | -19.4913
11.7324 | -34.5255 | -19.5001
11.7385 | -34.526 | -19.5088
11.7446 | -34.5263 | -19.5176
11.7507 | -34.5268 | -19.5259
11.7568 | -34.5272 | -19.5346
11.7629 | -34.5276 | -19.5433
11.769 | -34.528 | -19.5518
11.7751 | -34.5284 | -19.5605
11.7812 | -34.5288 | -19.5691
11.7873 | -34.5292 | -19.5777
11.7934 | -34.5296 | -19.5865
11.7995 | -34.53 | -19.5951
11.8056 | -34.5304 | -19.6038
11.8117 | -34.5308 | -19.6124
11.8178 | -34.5312 | -19.6207
11.824 | -34.5316 | -19.6292
11.8301 | -34.532 | -19.638
11.8362 | -34.5324 | -19.6467
11.8423 | -34.5327 | -19.6552
11.8484 | -34.5331 | -19.664
11.8545 | -34.5336 | -19.6724
11.8606 | -34.534 | -19.6809
11.8667 | -34.5344 | -19.6896
11.8728 | -34.5348 | -19.698
11.8789 | -34.5351 | -19.7066
11.885 | -34.5355 | -19.7154
11.8911 | -34.536 | -19.7237
11.8972 | -34.5364 | -19.7322
11.9033 | -34.5368 | -19.7409
11.9094 | -34.5372 | -19.7494
11.9155 | -34.5376 | -19.758
11.9216 | -34.538 | -19.7665
11.9277 | -34.5384 | -19.7751
11.9338 | -34.5387 | -19.7837
11.9399 | -34.5392 | -19.7924
11.946 | -34.5396 | -19.8009
11.9521 | -34.54 | -19.8093
11.9582 | -34.5404 | -19.8177
11.9644 | -34.5408 | -19.8262
11.9705 | -34.5412 | -19.8348
11.9766 | -34.5416 | -19.8434
11.9827 | -34.5421 | -19.8519
11.9888 | -34.5425 | -19.8603
11.9949 | -34.5429 | -19.8683
12.001 | -34.5433 | -19.8769
12.0071 | -34.5437 | -19.8852
12.0132 | -34.5441 | -19.8937
12.0193 | -34.5445 | -19.9022
12.0254 | -34.5449 | -19.9107
12.0315 | -34.5453 | -19.9192
12.0376 | -34.5457 | -19.9277
12.0437 | -34.5461 | -19.9361
12.0498 | -34.5465 | -19.9445
12.0559 | -34.5469 | -19.9529
12.062 | -34.5474 | -19.961
12.0681 | -34.5478 | -19.9695
12.0742 | -34.5482 | -19.9779
12.0803 | -34.5486 | -19.9867
12.0864 | -34.549 | -19.9953
12.0925 | -34.5495 | -20.0037
12.0986 | -34.5499 | -20.0121
12.1047 | -34.5503 | -20.0206
12.1109 | -34.5508 | -20.0288
12.117 | -34.5512 | -20.0373
12.1231 | -34.5516 | -20.0456
12.1292 | -34.552 | -20.054
12.1353 | -34.5524 | -20.0625
12.1414 | -34.5529 | -20.071
12.1475 | -34.5533 | -20.0796
12.1536 | -34.5537 | -20.088
12.1597 | -34.5541 | -20.0965
12.1658 | -34.5545 | -20.1048
12.1719 | -34.5549 | -20.1129
12.178 | -34.5553 | -20.1214
12.1841 | -34.5558 | -20.1299
12.1902 | -34.5562 | -20.1384
12.1963 | -34.5566 | -20.1469
12.2024 | -34.557 | -20.155
12.2085 | -34.5574 | -20.1634
12.2146 | -34.5579 | -20.172
12.2207 | -34.5583 | -20.1803
12.2268 | -34.5587 | -20.1886
12.2329 | -34.5591 | -20.197
12.239 | -34.5596 | -20.205
12.2451 | -34.56 | -20.2134
12.2513 | -34.5604 | -20.2219
12.2574 | -34.5609 | -20.2305
12.2635 | -34.5613 | -20.239
12.2696 | -34.5617 | -20.2471
12.2757 | -34.5621 | -20.2554
12.2818 | -34.5626 | -20.2639
12.2879 | -34.563 | -20.2722
12.294 | -34.5634 | -20.2803
12.3001 | -34.5639 | -20.2888
12.3062 | -34.5643 | -20.2972
12.3123 | -34.5648 | -20.3053
12.3184 | -34.5652 | -20.3136
12.3245 | -34.5656 | -20.3217
12.3306 | -34.566 | -20.3301
12.3367 | -34.5664 | -20.3384
12.3428 | -34.5669 | -20.3467
12.3489 | -34.5673 | -20.3547
12.355 | -34.5678 | -20.3632
12.3611 | -34.5682 | -20.3716
12.3672 | -34.5686 | -20.3798
12.3733 | -34.5691 | -20.3882
12.3794 | -34.5695 | -20.3961
12.3855 | -34.5699 | -20.4043
12.3916 | -34.5703 | -20.4127
12.3978 | -34.5708 | -20.4211
12.4039 | -34.5712 | -20.4294
12.41 | -34.5717 | -20.4378
12.4161 | -34.5721 | -20.4458
12.4222 | -34.5726 | -20.4541
12.4283 | -34.573 | -20.4623
12.4344 | -34.5735 | -20.4707
12.4405 | -34.5739 | -20.4791
12.4466 | -34.5743 | -20.4873
12.4527 | -34.5747 | -20.4955
12.4588 | -34.5752 | -20.5038
12.4649 | -34.5756 | -20.5122
12.471 | -34.5761 | -20.5207
12.4771 | -34.5765 | -20.5289
12.4832 | -34.5769 | -20.5371
12.4893 | -34.5774 | -20.5454
12.4954 | -34.5778 | -20.5535
12.5015 | -34.5782 | -20.5618
12.5076 | -34.5787 | -20.5702
12.5137 | -34.5791 | -20.5786
12.5198 | -34.5796 | -20.5868
12.5259 | -34.58 | -20.5951
12.532 | -34.5804 | -20.6034
12.5382 | -34.5809 | -20.6117
12.5443 | -34.5813 | -20.6199
12.5504 | -34.5817 | -20.6283
12.5565 | -34.5822 | -20.6365
12.5626 | -34.5827 | -20.6449
12.5687 | -34.5831 | -20.6533
12.5748 | -34.5836 | -20.6616
12.5809 | -34.584 | -20.6699
12.587 | -34.5844 | -20.6783
12.5931 | -34.5849 | -20.6865
12.5992 | -34.5854 | -20.6946
12.6053 | -34.5858 | -20.7029
12.6114 | -34.5863 | -20.7111
12.6175 | -34.5867 | -20.7194
12.6236 | -34.5871 | -20.7278
12.6297 | -34.5876 | -20.7361
12.6358 | -34.588 | -20.7441
12.6419 | -34.5885 | -20.7524
12.648 | -34.5889 | -20.7608
12.6541 | -34.5894 | -20.7691
12.6602 | -34.5898 | -20.7772
12.6663 | -34.5903 | -20.7854
12.6724 | -34.5907 | -20.7935
12.6785 | -34.5912 | -20.8016
12.6847 | -34.5916 | -20.8098
12.6908 | -34.5921 | -20.8181
12.6969 | -34.5926 | -20.8262
12.703 | -34.593 | -20.8347
12.7091 | -34.5934 | -20.843
12.7152 | -34.5939 | -20.8511
12.7213 | -34.5943 | -20.8595
12.7274 | -34.5948 | -20.8678
12.7335 | -34.5952 | -20.8759
12.7396 | -34.5957 | -20.8841
12.7457 | -34.5962 | -20.8925
12.7518 | -34.5966 | -20.9007
12.7579 | -34.5971 | -20.9092
12.764 | -34.5975 | -20.9175
12.7701 | -34.598 | -20.9259
12.7762 | -34.5984 | -20.934
12.7823 | -34.5989 | -20.9424
12.7884 | -34.5994 | -20.9506
12.7945 | -34.5998 | -20.9592
12.8006 | -34.6003 | -20.9674
12.8067 | -34.6007 | -20.9757
12.8128 | -34.6012 | -20.9838
12.8189 | -34.6016 | -20.992
12.8251 | -34.602 | -21.0003
12.8312 | -34.6025 | -21.0087
12.8373 | -34.603 | -21.0172
12.8434 | -34.6035 | -21.0255
12.8495 | -34.6039 | -21.0336
12.8556 | -34.6044 | -21.0418
12.8617 | -34.6048 | -21.05
12.8678 | -34.6053 | -21.0583
12.8739 | -34.6057 | -21.0665
12.88 | -34.6062 | -21.0749
12.8861 | -34.6066 | -21.083
12.8922 | -34.6071 | -21.0914
12.8983 | -34.6075 | -21.0996
12.9044 | -34.608 | -21.1077
12.9105 | -34.6085 | -21.1159
12.9166 | -34.6089 | -21.1241
12.9227 | -34.6094 | -21.1324
12.9288 | -34.6098 | -21.1407
12.9349 | -34.6103 | -21.149
12.941 | -34.6108 | -21.1573
12.9471 | -34.6112 | -21.1657
12.9532 | -34.6117 | -21.1739
12.9593 | -34.6122 | -21.1823
12.9654 | -34.6126 | -21.1906
12.9716 | -34.6131 | -21.1991
12.9777 | -34.6136 | -21.2077
12.9838 | -34.6141 | -21.2159
12.9899 | -34.6145 | -21.2241
12.996 | -34.615 | -21.2325
13.0021 | -34.6154 | -21.2409
13.0082 | -34.6159 | -21.249
13.0143 | -34.6164 | -21.2572
13.0204 | -34.6169 | -21.2654
13.0265 | -34.6174 | -21.2736
13.0326 | -34.6178 | -21.2819
13.0387 | -34.6183 | -21.2903
13.0448 | -34.6187 | -21.2986
13.0509 | -34.6192 | -21.3068
13.057 | -34.6197 | -21.315
13.0631 | -34.6201 | -21.3233
13.0692 | -34.6206 | -21.3315
13.0753 | -34.6211 | -21.3399
13.0814 | -34.6216 | -21.3483
13.0875 | -34.622 | -21.3567
13.0936 | -34.6225 | -21.3651
13.0997 | -34.6229 | -21.3733
13.1058 | -34.6234 | -21.3817
13.112 | -34.6239 | -21.3899
13.1181 | -34.6244 | -21.3984
13.1242 | -34.6249 | -21.4068
13.1303 | -34.6254 | -21.415
13.1364 | -34.6258 | -21.4233
13.1425 | -34.6262 | -21.4315
13.1486 | -34.6267 | -21.4399
13.1547 | -34.6272 | -21.4484
13.1608 | -34.6276 | -21.4567
13.1669 | -34.6281 | -21.465
13.173 | -34.6286 | -21.4734
13.1791 | -34.6291 | -21.4818
13.1852 | -34.6296 | -21.4902
13.1913 | -34.6301 | -21.4987
13.1974 | -34.6305 | -21.5072
13.2035 | -34.631 | -21.5156
13.2096 | -34.6315 | -21.5238
13.2157 | -34.6319 | -21.5322
13.2218 | -34.6324 | -21.5402
13.2279 | -34.6329 | -21.549
13.234 | -34.6334 | -21.5572
13.2401 | -34.6339 | -21.5655
13.2462 | -34.6343 | -21.5738
13.2524 | -34.6348 | -21.5822
13.2585 | -34.6353 | -21.5907
13.2646 | -34.6357 | -21.5992
13.2707 | -34.6362 | -21.6077
13.2768 | -34.6367 | -21.6162
13.2829 | -34.6372 | -21.6245
13.289 | -34.6377 | -21.6326
13.2951 | -34.6382 | -21.6412
13.3012 | -34.6386 | -21.6496
13.3073 | -34.6391 | -21.658
13.3134 | -34.6396 | -21.6664
13.3195 | -34.64 | -21.6748
13.3256 | -34.6404 | -21.6831
13.3317 | -34.6409 | -21.6918
13.3378 | -34.6413 | -21.7003
13.3439 | -34.6418 | -21.7087
13.35 | -34.6423 | -21.717
13.3561 | -34.6427 | -21.7254
13.3622 | -34.6432 | -21.7339
13.3683 | -34.6437 | -21.7425
13.3744 | -34.6442 | -21.7508
13.3805 | -34.6447 | -21.7593
13.3866 | -34.6451 | -21.7675
13.3927 | -34.6456 | -21.776
13.3989 | -34.6461 | -21.7847
13.405 | -34.6465 | -21.7932
13.4111 | -34.647 | -21.8015
13.4172 | -34.6474 | -21.8101
13.4233 | -34.6479 | -21.8187
13.4294 | -34.6483 | -21.8272
13.4355 | -34.6488 | -21.8358
13.4416 | -34.6493 | -21.8441
13.4477 | -34.6498 | -21.8525
13.4538 | -34.6502 | -21.8608
13.4599 | -34.6507 | -21.8694
13.466 | -34.6512 | -21.8781
13.4721 | -34.6517 | -21.8868
13.4782 | -34.6522 | -21.895
13.4843 | -34.6527 | -21.9036
13.4904 | -34.6532 | -21.9118
13.4965 | -34.6536 | -21.9203
13.5026 | -34.6541 | -21.9289
13.5087 | -34.6546 | -21.9374
13.5148 | -34.655 | -21.9459
13.5209 | -34.6555 | -21.9545
13.527 | -34.6559 | -21.9631
13.5331 | -34.6564 | -21.9717
13.5393 | -34.6569 | -21.9802
13.5454 | -34.6574 | -21.9883
13.5515 | -34.6578 | -21.9969
13.5576 | -34.6582 | -22.0054
13.5637 | -34.6587 | -22.0142
13.5698 | -34.6592 | -22.0227
13.5759 | -34.6597 | -22.031
13.582 | -34.6601 | -22.0394
13.5881 | -34.6606 | -22.0479
13.5942 | -34.6611 | -22.0566
13.6003 | -34.6616 | -22.065
13.6064 | -34.6621 | -22.0736
13.6125 | -34.6626 | -22.0824
13.6186 | -34.6631 | -22.0909
13.6247 | -34.6636 | -22.0994
13.6308 | -34.664 | -22.1079
13.6369 | -34.6645 | -22.1165
13.643 | -34.665 | -22.1252
13.6491 | -34.6654 | -22.1337
13.6552 | -34.6659 | -22.1422
13.6613 | -34.6664 | -22.1507
13.6674 | -34.6668 | -22.1591
13.6735 | -34.6673 | -22.1678
13.6796 | -34.6678 | -22.1763
13.6858 | -34.6682 | -22.1848
13.6919 | -34.6687 | -22.1934
13.698 | -34.6692 | -22.2019
13.7041 | -34.6697 | -22.2104
13.7102 | -34.6702 | -22.2188
13.7163 | -34.6707 | -22.2275
13.7224 | -34.6711 | -22.2361
13.7285 | -34.6716 | -22.2447
13.7346 | -34.6721 | -22.2532
13.7407 | -34.6726 | -22.2616
13.7468 | -34.673 | -22.2702
13.7529 | -34.6735 | -22.2788
13.759 | -34.674 | -22.2874
13.7651 | -34.6744 | -22.2959
13.7712 | -34.6749 | -22.3045
13.7773 | -34.6754 | -22.3131
13.7834 | -34.6759 | -22.3217
13.7895 | -34.6763 | -22.3303
13.7956 | -34.6768 | -22.3389
13.8017 | -34.6773 | -22.3473
13.8078 | -34.6778 | -22.3558
13.8139 | -34.6783 | -22.3642
13.82 | -34.6787 | -22.3726
13.8262 | -34.6792 | -22.3813
13.8323 | -34.6797 | -22.3899
13.8384 | -34.6802 | -22.3984
13.8445 | -34.6807 | -22.4069
13.8506 | -34.6812 | -22.4155
13.8567 | -34.6816 | -22.4242
13.8628 | -34.6821 | -22.4326
13.8689 | -34.6826 | -22.4413
13.875 | -34.6831 | -22.4497
13.8811 | -34.6835 | -22.4583
13.8872 | -34.684 | -22.4669
13.8933 | -34.6845 | -22.4753
13.8994 | -34.6849 | -22.484
13.9055 | -34.6854 | -22.4926
13.9116 | -34.6859 | -22.5011
13.9177 | -34.6864 | -22.5097
13.9238 | -34.6868 | -22.5182
13.9299 | -34.6873 | -22.5268
13.936 | -34.6877 | -22.5353
13.9421 | -34.6882 | -22.5437
13.9482 | -34.6887 | -22.5522
13.9543 | -34.6891 | -22.5609
13.9604 | -34.6896 | -22.5693
13.9665 | -34.6901 | -22.5779
13.9727 | -34.6906 | -22.5866
13.9788 | -34.6911 | -22.5952
13.9849 | -34.6915 | -22.6038
13.991 | -34.692 | -22.6124
13.9971 | -34.6925 | -22.6208
14.0032 | -34.6929 | -22.6294
14.0093 | -34.6934 | -22.638
14.0154 | -34.6938 | -22.6464
14.0215 | -34.6943 | -22.6548
14.0276 | -34.6948 | -22.6634
14.0337 | -34.6953 | -22.672
14.0398 | -34.6958 | -22.6807
14.0459 | -34.6963 | -22.6891
14.052 | -34.6967 | -22.6979
14.0581 | -34.6972 | -22.7065
14.0642 | -34.6977 | -22.7152
14.0703 | -34.6982 | -22.7239
14.0764 | -34.6986 | -22.7324
14.0825 | -34.6991 | -22.741
14.0886 | -34.6996 | -22.7495
14.0947 | -34.7 | -22.7581
14.1008 | -34.7005 | -22.7667
14.1069 | -34.701 | -22.7751
14.1131 | -34.7014 | -22.7836
14.1192 | -34.7019 | -22.792
14.1253 | -34.7024 | -22.8005
14.1314 | -34.7029 | -22.8091
14.1375 | -34.7033 | -22.8177
14.1436 | -34.7038 | -22.8263
14.1497 | -34.7043 | -22.8347
14.1558 | -34.7048 | -22.8432
14.1619 | -34.7053 | -22.8517
14.168 | -34.7057 | -22.8603
14.1741 | -34.7062 | -22.8689
14.1802 | -34.7067 | -22.8774
14.1863 | -34.7072 | -22.886
14.1924 | -34.7076 | -22.8946
14.1985 | -34.7081 | -22.9031
14.2046 | -34.7086 | -22.9118
14.2107 | -34.7091 | -22.92
14.2168 | -34.7095 | -22.9287
14.2229 | -34.71 | -22.9372
14.229 | -34.7105 | -22.9459
14.2351 | -34.7109 | -22.9544
14.2412 | -34.7114 | -22.9629
14.2473 | -34.7119 | -22.9713
14.2534 | -34.7123 | -22.9797
14.2596 | -34.7128 | -22.9881
14.2657 | -34.7132 | -22.9967
14.2718 | -34.7137 | -23.0053
14.2779 | -34.7142 | -23.0139
14.284 | -34.7147 | -23.0224
14.2901 | -34.7152 | -23.0309
14.2962 | -34.7157 | -23.0394
14.3023 | -34.7161 | -23.048
14.3084 | -34.7166 | -23.0565
14.3145 | -34.7171 | -23.0649
14.3206 | -34.7176 | -23.0738
14.3267 | -34.7181 | -23.0823
14.3328 | -34.7186 | -23.0909
14.3389 | -34.719 | -23.0994
14.345 | -34.7195 | -23.1079
14.3511 | -34.72 | -23.1165
14.3572 | -34.7205 | -23.125
14.3633 | -34.7209 | -23.1337
14.3694 | -34.7214 | -23.1421
14.3755 | -34.7219 | -23.1506
14.3816 | -34.7224 | -23.1593
14.3877 | -34.7229 | -23.1677
14.3938 | -34.7233 | -23.1765
14.4 | -34.7238 | -23.1849
14.4061 | -34.7243 | -23.1934
14.4122 | -34.7247 | -23.2021
14.4183 | -34.7252 | -23.2106
14.4244 | -34.7257 | -23.2191
14.4305 | -34.7262 | -23.2276
14.4366 | -34.7267 | -23.2361
14.4427 | -34.7271 | -23.2448
14.4488 | -34.7276 | -23.2533
14.4549 | -34.7281 | -23.2618
14.461 | -34.7286 | -23.2702
14.4671 | -34.729 | -23.2788
14.4732 | -34.7295 | -23.2875
14.4793 | -34.73 | -23.2961
14.4854 | -34.7304 | -23.3044
14.4915 | -34.7309 | -23.3127
14.4976 | -34.7314 | -23.3211
14.5037 | -34.7319 | -23.3296
14.5098 | -34.7323 | -23.3381
14.5159 | -34.7328 | -23.3466
14.522 | -34.7332 | -23.355
14.5281 | -34.7337 | -23.3634
14.5342 | -34.7342 | -23.372
14.5403 | -34.7347 | -23.3803
14.5465 | -34.7351 | -23.3889
14.5526 | -34.7356 | -23.3975
14.5587 | -34.7361 | -23.406
14.5648 | -34.7366 | -23.4143
14.5709 | -34.737 | -23.4228
14.577 | -34.7375 | -23.4312
14.5831 | -34.738 | -23.4396
14.5892 | -34.7385 | -23.448
14.5953 | -34.7389 | -23.4564
14.6014 | -34.7394 | -23.465
14.6075 | -34.7399 | -23.4733
14.6136 | -34.7403 | -23.482
14.6197 | -34.7408 | -23.4904
14.6258 | -34.7412 | -23.4989
14.6319 | -34.7417 | -23.5072
14.638 | -34.7422 | -23.5157
14.6441 | -34.7426 | -23.5242
14.6502 | -34.7431 | -23.5328
14.6563 | -34.7436 | -23.5412
14.6624 | -34.7441 | -23.5498
14.6685 | -34.7445 | -23.5579
14.6746 | -34.745 | -23.5664
14.6807 | -34.7455 | -23.5748
14.6869 | -34.7459 | -23.5831
14.693 | -34.7464 | -23.5915
14.6991 | -34.7469 | -23.5999
14.7052 | -34.7474 | -23.6081
14.7113 | -34.7479 | -23.6166
14.7174 | -34.7483 | -23.6252
14.7235 | -34.7488 | -23.6335
14.7296 | -34.7493 | -23.6418
14.7357 | -34.7498 | -23.65
14.7418 | -34.7502 | -23.6583
14.7479 | -34.7507 | -23.6669
14.754 | -34.7512 | -23.6753
14.7601 | -34.7517 | -23.6837
14.7662 | -34.7522 | -23.6921
14.7723 | -34.7526 | -23.7003
14.7784 | -34.7531 | -23.7087
14.7845 | -34.7535 | -23.7169
14.7906 | -34.754 | -23.7255
14.7967 | -34.7545 | -23.734
14.8028 | -34.7549 | -23.7423
14.8089 | -34.7554 | -23.7506
14.815 | -34.7559 | -23.7589
14.8211 | -34.7564 | -23.7672
14.8272 | -34.7569 | -23.7758
14.8334 | -34.7573 | -23.7841
14.8395 | -34.7578 | -23.7924
14.8456 | -34.7583 | -23.8007
14.8517 | -34.7588 | -23.809
14.8578 | -34.7592 | -23.8176
14.8639 | -34.7597 | -23.8261
14.87 | -34.7602 | -23.8344
14.8761 | -34.7607 | -23.8426
14.8822 | -34.7612 | -23.8511
14.8883 | -34.7616 | -23.8594
14.8944 | -34.7621 | -23.8677
14.9005 | -34.7625 | -23.876
14.9066 | -34.763 | -23.8843
14.9127 | -34.7635 | -23.8926
14.9188 | -34.764 | -23.9007
14.9249 | -34.7644 | -23.9091
14.931 | -34.7649 | -23.9175
14.9371 | -34.7654 | -23.9258
14.9432 | -34.7659 | -23.9342
14.9493 | -34.7664 | -23.9423
14.9554 | -34.7668 | -23.9506
14.9615 | -34.7673 | -23.9587
14.9676 | -34.7678 | -23.9672
14.9738 | -34.7683 | -23.9754
14.9799 | -34.7687 | -23.9838
14.986 | -34.7692 | -23.9919
14.9921 | -34.7697 | -24.0001
14.9982 | -34.7701 | -24.0082
15.0043 | -34.7706 | -24.0165
15.0104 | -34.771 | -24.0249
15.0165 | -34.7715 | -24.0332
15.0226 | -34.772 | -24.0412
15.0287 | -34.7725 | -24.0495
15.0348 | -34.773 | -24.0577
15.0409 | -34.7734 | -24.066
15.047 | -34.7739 | -24.0745
15.0531 | -34.7744 | -24.0826
15.0592 | -34.7749 | -24.091
15.0653 | -34.7754 | -24.0994
15.0714 | -34.7759 | -24.1076
15.0775 | -34.7764 | -24.1157
15.0836 | -34.7768 | -24.1241
15.0897 | -34.7773 | -24.1319
15.0958 | -34.7778 | -24.1404
15.1019 | -34.7783 | -24.1485
15.108 | -34.7788 | -24.1566
15.1141 | -34.7793 | -24.1649
15.1203 | -34.7797 | -24.1733
15.1264 | -34.7802 | -24.1814
15.1325 | -34.7807 | -24.1897
15.1386 | -34.7812 | -24.1978
15.1447 | -34.7817 | -24.2059
15.1508 | -34.7822 | -24.2139
15.1569 | -34.7826 | -24.222
15.163 | -34.7831 | -24.2299
15.1691 | -34.7836 | -24.2382
15.1752 | -34.7841 | -24.2465
15.1813 | -34.7845 | -24.2545
15.1874 | -34.785 | -24.2628
15.1935 | -34.7855 | -24.2708
15.1996 | -34.786 | -24.2789
15.2057 | -34.7865 | -24.287
15.2118 | -34.787 | -24.2953
15.2179 | -34.7875 | -24.3035
15.224 | -34.788 | -24.3116
15.2301 | -34.7885 | -24.3195
15.2362 | -34.789 | -24.3276
15.2423 | -34.7895 | -24.3359
15.2484 | -34.79 | -24.3441
15.2545 | -34.7904 | -24.3522
15.2607 | -34.7909 | -24.3603
15.2668 | -34.7914 | -24.3681
15.2729 | -34.7919 | -24.3761
15.279 | -34.7924 | -24.3845
15.2851 | -34.7929 | -24.3925
15.2912 | -34.7934 | -24.4007
15.2973 | -34.7938 | -24.4091
15.3034 | -34.7943 | -24.417
15.3095 | -34.7948 | -24.4249
15.3156 | -34.7953 | -24.4333
15.3217 | -34.7957 | -24.4412
15.3278 | -34.7962 | -24.4491
15.3339 | -34.7967 | -24.457
15.34 | -34.7971 | -24.4647
15.3461 | -34.7976 | -24.4729
15.3522 | -34.798 | -24.481
15.3583 | -34.7985 | -24.4893
15.3644 | -34.799 | -24.4973
15.3705 | -34.7995 | -24.505
15.3766 | -34.7999 | -24.5134
15.3827 | -34.8004 | -24.5214
15.3888 | -34.8009 | -24.53
15.3949 | -34.8014 | -24.5378
15.401 | -34.8018 | -24.5457
15.4072 | -34.8023 | -24.5535
15.4133 | -34.8028 | -24.5618
15.4194 | -34.8033 | -24.5697
15.4255 | -34.8037 | -24.5777
15.4316 | -34.8042 | -24.5857
15.4377 | -34.8046 | -24.5939
15.4438 | -34.8051 | -24.6017
15.4499 | -34.8056 | -24.6097
15.456 | -34.8061 | -24.6178
15.4621 | -34.8066 | -24.6257
15.4682 | -34.8071 | -24.634
15.4743 | -34.8076 | -24.6417
15.4804 | -34.8081 | -24.6496
15.4865 | -34.8085 | -24.6575
15.4926 | -34.8089 | -24.6653
15.4987 | -34.8095 | -24.6738
15.5048 | -34.81 | -24.6819
15.5109 | -34.8105 | -24.6898
15.517 | -34.8109 | -24.6978
15.5231 | -34.8114 | -24.7057
15.5292 | -34.8119 | -24.7142
15.5353 | -34.8124 | -24.7222
15.5414 | -34.8128 | -24.7301
15.5476 | -34.8133 | -24.7376
15.5537 | -34.8138 | -24.7454
15.5598 | -34.8142 | -24.7537
15.5659 | -34.8148 | -24.7618
15.572 | -34.8153 | -24.7703
15.5781 | -34.8157 | -24.7782
15.5842 | -34.8162 | -24.7861
15.5903 | -34.8167 | -24.7945
15.5964 | -34.8172 | -24.8021
15.6025 | -34.8177 | -24.8103
15.6086 | -34.8183 | -24.8183
15.6147 | -34.8188 | -24.8266
15.6208 | -34.8193 | -24.8343
15.6269 | -34.8199 | -24.8423
15.633 | -34.8203 | -24.8506
15.6391 | -34.8207 | -24.8586
15.6452 | -34.8212 | -24.8669
15.6513 | -34.8217 | -24.8746
15.6574 | -34.8223 | -24.8822
15.6635 | -34.8228 | -24.8902
15.6696 | -34.8233 | -24.8981
15.6757 | -34.8238 | -24.9059
15.6818 | -34.8242 | -24.9141
15.688 | -34.8247 | -24.9218
15.6941 | -34.8251 | -24.93
15.7002 | -34.8256 | -24.9379
15.7063 | -34.8261 | -24.946
15.7124 | -34.8265 | -24.9542
15.7185 | -34.827 | -24.9622
15.7246 | -34.8275 | -24.9699
15.7307 | -34.828 | -24.9777
15.7368 | -34.8286 | -24.9856
15.7429 | -34.8291 | -24.9935
15.749 | -34.8296 | -25.0016
15.7551 | -34.8301 | -25.0094
15.7612 | -34.8306 | -25.017
15.7673 | -34.8311 | -25.0251
15.7734 | -34.8316 | -25.0328
15.7795 | -34.8321 | -25.0406
15.7856 | -34.8326 | -25.0485
15.7917 | -34.8331 | -25.0561
15.7978 | -34.8336 | -25.0646
15.8039 | -34.834 | -25.0725
15.81 | -34.8345 | -25.0803
15.8161 | -34.835 | -25.0881
15.8222 | -34.8355 | -25.0959
15.8283 | -34.836 | -25.1034
15.8345 | -34.8365 | -25.1112
15.8406 | -34.837 | -25.1191
15.8467 | -34.8375 | -25.1269
15.8528 | -34.838 | -25.135
15.8589 | -34.8385 | -25.1431
15.865 | -34.839 | -25.1506
15.8711 | -34.8395 | -25.1584
15.8772 | -34.8401 | -25.1659
15.8833 | -34.8405 | -25.1734
15.8894 | -34.841 | -25.1817
15.8955 | -34.8414 | -25.1897
15.9016 | -34.8419 | -25.1972
15.9077 | -34.8424 | -25.205
15.9138 | -34.8428 | -25.2128
15.9199 | -34.8433 | -25.2206
15.926 | -34.8438 | -25.2282
15.9321 | -34.8443 | -25.2359
15.9382 | -34.8447 | -25.2439
15.9443 | -34.8452 | -25.2515
15.9504 | -34.8457 | -25.2594
15.9565 | -34.8463 | -25.2671
15.9626 | -34.8467 | -25.2757
15.9687 | -34.8472 | -25.2831
15.9749 | -34.8478 | -25.2909
15.981 | -34.8482 | -25.2987
15.9871 | -34.8487 | -25.3065
15.9932 | -34.8492 | -25.314
15.9993 | -34.8496 | -25.3217
16.0054 | -34.8501 | -25.3291
16.0115 | -34.8506 | -25.3372
16.0176 | -34.8511 | -25.345
16.0237 | -34.8516 | -25.3532
16.0298 | -34.8521 | -25.3611
16.0359 | -34.8525 | -25.3688
16.042 | -34.853 | -25.3762
16.0481 | -34.8535 | -25.3842
16.0542 | -34.854 | -25.392
16.0603 | -34.8544 | -25.3994
16.0664 | -34.8549 | -25.4072
16.0725 | -34.8554 | -25.4147
16.0786 | -34.8559 | -25.4224
16.0847 | -34.8564 | -25.4305
16.0908 | -34.8569 | -25.4384
16.0969 | -34.8573 | -25.4463
16.103 | -34.8578 | -25.4542
16.1091 | -34.8583 | -25.4617
16.1152 | -34.8588 | -25.4695
16.1214 | -34.8593 | -25.4772
16.1275 | -34.8598 | -25.485
16.1336 | -34.8603 | -25.4924
16.1397 | -34.8608 | -25.5001
16.1458 | -34.8613 | -25.5076
16.1519 | -34.8617 | -25.5157
16.158 | -34.8622 | -25.5234
16.1641 | -34.8627 | -25.531
16.1702 | -34.8631 | -25.5388
16.1763 | -34.8636 | -25.5467
16.1824 | -34.8641 | -25.5541
16.1885 | -34.8646 | -25.5617
16.1946 | -34.865 | -25.5694
16.2007 | -34.8655 | -25.5769
16.2068 | -34.866 | -25.5845
16.2129 | -34.8665 | -25.5922
16.219 | -34.867 | -25.5996
16.2251 | -34.8675 | -25.6069
16.2312 | -34.868 | -25.6147
16.2373 | -34.8684 | -25.6223
16.2434 | -34.8689 | -25.63
16.2495 | -34.8694 | -25.6373
16.2556 | -34.8699 | -25.6454
16.2618 | -34.8703 | -25.6529
16.2679 | -34.8708 | -25.6608
16.274 | -34.8713 | -25.6688
16.2801 | -34.8718 | -25.6764
16.2862 | -34.8723 | -25.6839
16.2923 | -34.8728 | -25.6915
16.2984 | -34.8733 | -25.6991
16.3045 | -34.8738 | -25.7068
16.3106 | -34.8743 | -25.7146
16.3167 | -34.8748 | -25.7223
16.3228 | -34.8752 | -25.7298
16.3289 | -34.8757 | -25.7374
16.335 | -34.8762 | -25.7449
16.3411 | -34.8767 | -25.7525
16.3472 | -34.8772 | -25.7602
16.3533 | -34.8776 | -25.7679
16.3594 | -34.8781 | -25.7752
16.3655 | -34.8786 | -25.7829
16.3716 | -34.8791 | -25.7904
16.3777 | -34.8796 | -25.798
16.3838 | -34.8801 | -25.8056
16.3899 | -34.8805 | -25.813
16.396 | -34.881 | -25.8207
16.4021 | -34.8815 | -25.8283
16.4083 | -34.882 | -25.8357
16.4144 | -34.8825 | -25.8435
16.4205 | -34.883 | -25.851
16.4266 | -34.8835 | -25.8584
16.4327 | -34.884 | -25.8658
16.4388 | -34.8845 | -25.8735
16.4449 | -34.885 | -25.8809
16.451 | -34.8855 | -25.8885
16.4571 | -34.886 | -25.8964
16.4632 | -34.8865 | -25.9038
16.4693 | -34.8869 | -25.9114
16.4754 | -34.8874 | -25.919
16.4815 | -34.8879 | -25.9268
16.4876 | -34.8883 | -25.9342
16.4937 | -34.8888 | -25.9417
16.4998 | -34.8893 | -25.9491
16.5059 | -34.8898 | -25.9566
16.512 | -34.8903 | -25.9642
16.5181 | -34.8908 | -25.972
16.5242 | -34.8913 | -25.9796
16.5303 | -34.8917 | -25.987
16.5364 | -34.8922 | -25.9945
16.5425 | -34.8927 | -26.0022
16.5487 | -34.8932 | -26.0098
16.5548 | -34.8936 | -26.0173
16.5609 | -34.8941 | -26.025
16.567 | -34.8946 | -26.0323
16.5731 | -34.8951 | -26.04
16.5792 | -34.8956 | -26.0478
16.5853 | -34.8961 | -26.0554
16.5914 | -34.8966 | -26.0631
16.5975 | -34.8971 | -26.0707
16.6036 | -34.8976 | -26.0781
16.6097 | -34.898 | -26.0855
16.6158 | -34.8985 | -26.0929
16.6219 | -34.899 | -26.1004
16.628 | -34.8995 | -26.1079
16.6341 | -34.9 | -26.1153
16.6402 | -34.9005 | -26.1227
16.6463 | -34.9009 | -26.1302
16.6524 | -34.9014 | -26.1376
16.6585 | -34.9019 | -26.1454
16.6646 | -34.9024 | -26.1527
16.6707 | -34.9028 | -26.1601
16.6768 | -34.9033 | -26.1678
16.6829 | -34.9038 | -26.1753
16.689 | -34.9043 | -26.1828
16.6952 | -34.9047 | -26.1904
16.7013 | -34.9052 | -26.1978
16.7074 | -34.9057 | -26.205
16.7135 | -34.9061 | -26.2125
16.7196 | -34.9066 | -26.2202
16.7257 | -34.9071 | -26.2278
16.7318 | -34.9076 | -26.2352
16.7379 | -34.9081 | -26.2426
16.744 | -34.9085 | -26.2499
16.7501 | -34.909 | -26.2574
16.7562 | -34.9096 | -26.265
16.7623 | -34.91 | -26.2724
16.7684 | -34.9105 | -26.28
16.7745 | -34.911 | -26.2876
16.7806 | -34.9115 | -26.2947
16.7867 | -34.912 | -26.3024
16.7928 | -34.9124 | -26.31
16.7989 | -34.9129 | -26.3174
16.805 | -34.9134 | -26.3252
16.8111 | -34.9139 | -26.3326
16.8172 | -34.9144 | -26.3401
16.8233 | -34.9149 | -26.3476
16.8294 | -34.9154 | -26.3552
16.8356 | -34.9158 | -26.3628
16.8417 | -34.9163 | -26.3701
16.8478 | -34.9168 | -26.3776
16.8539 | -34.9173 | -26.3851
16.86 | -34.9177 | -26.3928
16.8661 | -34.9182 | -26.4002
16.8722 | -34.9187 | -26.408
16.8783 | -34.9192 | -26.4154
16.8844 | -34.9197 | -26.4228
16.8905 | -34.9202 | -26.4303
16.8966 | -34.9206 | -26.4376
16.9027 | -34.9211 | -26.4452
16.9088 | -34.9216 | -26.4527
16.9149 | -34.922 | -26.4602
16.921 | -34.9226 | -26.4676
16.9271 | -34.923 | -26.4752
16.9332 | -34.9235 | -26.4826
16.9393 | -34.924 | -26.4902
16.9454 | -34.9244 | -26.4974
16.9515 | -34.9249 | -26.5048
16.9576 | -34.9254 | -26.5121
16.9637 | -34.926 | -26.5195
16.9698 | -34.9264 | -26.527
16.9759 | -34.9269 | -26.5343
16.9821 | -34.9273 | -26.5421
16.9882 | -34.9278 | -26.5496
16.9943 | -34.9283 | -26.5569
17.0004 | -34.9288 | -26.5644
17.0065 | -34.9293 | -26.5718
17.0126 | -34.9298 | -26.5793
17.0187 | -34.9303 | -26.5869
17.0248 | -34.9308 | -26.5942
17.0309 | -34.9312 | -26.6016
17.037 | -34.9317 | -26.6092
17.0431 | -34.9322 | -26.6167
17.0492 | -34.9327 | -26.6243
17.0553 | -34.9332 | -26.6316
17.0614 | -34.9337 | -26.6391
17.0675 | -34.9341 | -26.6468
17.0736 | -34.9346 | -26.6543
17.0797 | -34.935 | -26.6617
17.0858 | -34.9355 | -26.6694
17.0919 | -34.936 | -26.6769
17.098 | -34.9365 | -26.6841
17.1041 | -34.937 | -26.6913
17.1102 | -34.9374 | -26.6986
17.1163 | -34.9379 | -26.7061
17.1225 | -34.9384 | -26.7135
17.1286 | -34.9389 | -26.7208
17.1347 | -34.9393 | -26.7282
17.1408 | -34.9398 | -26.7359
17.1469 | -34.9403 | -26.7433
17.153 | -34.9408 | -26.7507
17.1591 | -34.9413 | -26.7582
17.1652 | -34.9418 | -26.7656
17.1713 | -34.9422 | -26.7729
17.1774 | -34.9427 | -26.7801
17.1835 | -34.9432 | -26.7875
17.1896 | -34.9436 | -26.7949
17.1957 | -34.9441 | -26.8022
17.2018 | -34.9446 | -26.8094
17.2079 | -34.9451 | -26.8169
17.214 | -34.9455 | -26.8245
17.2201 | -34.946 | -26.8319
17.2262 | -34.9465 | -26.8394
17.2323 | -34.9469 | -26.8467
17.2384 | -34.9474 | -26.8541
17.2445 | -34.9479 | -26.8613
17.2506 | -34.9484 | -26.8687
17.2567 | -34.9488 | -26.8762
17.2628 | -34.9493 | -26.8836
17.269 | -34.9498 | -26.8909
17.2751 | -34.9503 | -26.8983
17.2812 | -34.9507 | -26.9058
17.2873 | -34.9512 | -26.9132
17.2934 | -34.9517 | -26.9208
17.2995 | -34.9522 | -26.9284
17.3056 | -34.9527 | -26.9357
17.3117 | -34.9531 | -26.9431
17.3178 | -34.9536 | -26.9504
17.3239 | -34.954 | -26.9576
17.33 | -34.9545 | -26.965
17.3361 | -34.955 | -26.9723
17.3422 | -34.9554 | -26.9797
17.3483 | -34.9559 | -26.9872
17.3544 | -34.9564 | -26.9946
17.3605 | -34.9568 | -27.0019
17.3666 | -34.9573 | -27.0093
17.3727 | -34.9578 | -27.0167
17.3788 | -34.9583 | -27.024
17.3849 | -34.9588 | -27.0314
17.391 | -34.9592 | -27.0388
17.3971 | -34.9597 | -27.0461
17.4032 | -34.9601 | -27.0535
17.4094 | -34.9606 | -27.0609
17.4155 | -34.9611 | -27.0682
17.4216 | -34.9616 | -27.0754
17.4277 | -34.9621 | -27.0829
17.4338 | -34.9625 | -27.0904
17.4399 | -34.963 | -27.0977
17.446 | -34.9635 | -27.1049
17.4521 | -34.9639 | -27.1123
17.4582 | -34.9644 | -27.1195
17.4643 | -34.9648 | -27.127
17.4704 | -34.9653 | -27.1345
17.4765 | -34.9658 | -27.142
17.4826 | -34.9663 | -27.1492
17.4887 | -34.9667 | -27.1566
17.4948 | -34.9672 | -27.1639
17.5009 | -34.9677 | -27.1713
17.507 | -34.9681 | -27.1788
17.5131 | -34.9686 | -27.1862
17.5192 | -34.9691 | -27.1935
17.5253 | -34.9696 | -27.201
17.5314 | -34.97 | -27.2084
17.5375 | -34.9705 | -27.2157
17.5436 | -34.9709 | -27.2232
17.5497 | -34.9714 | -27.2305
17.5559 | -34.9719 | -27.2377
17.562 | -34.9724 | -27.245
17.5681 | -34.9729 | -27.2524
17.5742 | -34.9733 | -27.2598
17.5803 | -34.9738 | -27.2671
17.5864 | -34.9742 | -27.2744
17.5925 | -34.9747 | -27.2819
17.5986 | -34.9752 | -27.2891
17.6047 | -34.9757 | -27.2966
17.6108 | -34.9762 | -27.3039
17.6169 | -34.9766 | -27.3112
17.623 | -34.9771 | -27.3184
17.6291 | -34.9776 | -27.3257
17.6352 | -34.9781 | -27.3332
17.6413 | -34.9786 | -27.3405
17.6474 | -34.979 | -27.3479
17.6535 | -34.9795 | -27.3553
17.6596 | -34.9799 | -27.3625
17.6657 | -34.9804 | -27.37
17.6718 | -34.9809 | -27.3774
17.6779 | -34.9813 | -27.3848
17.684 | -34.9818 | -27.3922
17.6901 | -34.9823 | -27.3995
17.6963 | -34.9827 | -27.4068
17.7024 | -34.9832 | -27.4141
17.7085 | -34.9837 | -27.4213
17.7146 | -34.9841 | -27.4284
17.7207 | -34.9846 | -27.4356
17.7268 | -34.9851 | -27.4429
17.7329 | -34.9856 | -27.4502
17.739 | -34.986 | -27.4576
17.7451 | -34.9865 | -27.4648
17.7512 | -34.987 | -27.4722
17.7573 | -34.9875 | -27.4796
17.7634 | -34.9879 | -27.4868
17.7695 | -34.9884 | -27.494
17.7756 | -34.9889 | -27.5014
17.7817 | -34.9894 | -27.5087
17.7878 | -34.9898 | -27.5164
17.7939 | -34.9903 | -27.5236
17.8 | -34.9908 | -27.5308
17.8061 | -34.9912 | -27.538
17.8122 | -34.9917 | -27.5452
17.8183 | -34.9922 | -27.5526
17.8244 | -34.9926 | -27.56
17.8305 | -34.9931 | -27.5673
17.8366 | -34.9936 | -27.5744
17.8428 | -34.994 | -27.5817
17.8489 | -34.9945 | -27.5891
17.855 | -34.995 | -27.5964
17.8611 | -34.9955 | -27.6037
17.8672 | -34.996 | -27.6112
17.8733 | -34.9964 | -27.6187
17.8794 | -34.9969 | -27.626
17.8855 | -34.9974 | -27.6334
17.8916 | -34.9978 | -27.6408
17.8977 | -34.9983 | -27.6481
17.9038 | -34.9988 | -27.6553
17.9099 | -34.9992 | -27.6627
17.916 | -34.9997 | -27.67
17.9221 | -35.0002 | -27.6773
17.9282 | -35.0006 | -27.6845
17.9343 | -35.0011 | -27.6917
17.9404 | -35.0016 | -27.6991
17.9465 | -35.0021 | -27.7064
17.9526 | -35.0025 | -27.7138
17.9587 | -35.0029 | -27.7212
17.9648 | -35.0034 | -27.7284
17.9709 | -35.0038 | -27.7356
17.977 | -35.0043 | -27.7429
17.9832 | -35.0048 | -27.7501
17.9893 | -35.0052 | -27.7574
17.9954 | -35.0057 | -27.7648
18.0015 | -35.0062 | -27.7721
18.0076 | -35.0066 | -27.7793
18.0137 | -35.0071 | -27.7867
18.0198 | -35.0075 | -27.794
18.0259 | -35.008 | -27.8014
18.032 | -35.0084 | -27.8087
18.0381 | -35.0089 | -27.8161
18.0442 | -35.0094 | -27.8233
18.0503 | -35.0098 | -27.8307
18.0564 | -35.0103 | -27.838
18.0625 | -35.0108 | -27.8453
18.0686 | -35.0112 | -27.8525
18.0747 | -35.0117 | -27.8599
18.0808 | -35.0121 | -27.8671
18.0869 | -35.0126 | -27.8745
18.093 | -35.0131 | -27.8818
18.0991 | -35.0135 | -27.8892
18.1052 | -35.014 | -27.8965
18.1113 | -35.0144 | -27.9037
18.1174 | -35.0148 | -27.9107
18.1236 | -35.0153 | -27.918
18.1297 | -35.0158 | -27.9253
18.1358 | -35.0162 | -27.9327
18.1419 | -35.0167 | -27.9399
18.148 | -35.0172 | -27.9471
18.1541 | -35.0177 | -27.9543
18.1602 | -35.0181 | -27.9615
18.1663 | -35.0186 | -27.9687
18.1724 | -35.019 | -27.976
18.1785 | -35.0195 | -27.983
18.1846 | -35.0199 | -27.9902
18.1907 | -35.0204 | -27.9975
18.1968 | -35.0209 | -28.0048
18.2029 | -35.0213 | -28.0122
18.209 | -35.0218 | -28.0195
18.2151 | -35.0223 | -28.0269
18.2212 | -35.0227 | -28.0339
18.2273 | -35.0232 | -28.0413
18.2334 | -35.0236 | -28.0486
18.2395 | -35.0241 | -28.0556
18.2456 | -35.0246 | -28.063
18.2517 | -35.025 | -28.0703
18.2578 | -35.0255 | -28.0774
18.2639 | -35.026 | -28.0847
18.2701 | -35.0265 | -28.0919
18.2762 | -35.0269 | -28.0992
18.2823 | -35.0274 | -28.1067
18.2884 | -35.0279 | -28.1137
18.2945 | -35.0283 | -28.1208
18.3006 | -35.0288 | -28.1282
18.3067 | -35.0293 | -28.1353
18.3128 | -35.0298 | -28.1428
18.3189 | -35.0302 | -28.15
18.325 | -35.0307 | -28.157
18.3311 | -35.0311 | -28.1643
18.3372 | -35.0316 | -28.1716
18.3433 | -35.0321 | -28.1789
18.3494 | -35.0325 | -28.1861
18.3555 | -35.033 | -28.1934
18.3616 | -35.0335 | -28.2004
18.3677 | -35.034 | -28.2077
18.3738 | -35.0344 | -28.2149
18.3799 | -35.0349 | -28.2222
18.386 | -35.0354 | -28.2295
18.3921 | -35.0358 | -28.2367
18.3982 | -35.0363 | -28.2437
18.4043 | -35.0367 | -28.2508
18.4105 | -35.0372 | -28.2582
18.4166 | -35.0377 | -28.2654
18.4227 | -35.0381 | -28.2727
18.4288 | -35.0386 | -28.2798
18.4349 | -35.039 | -28.2871
18.441 | -35.0395 | -28.2943
18.4471 | -35.0399 | -28.3013
18.4532 | -35.0404 | -28.3085
18.4593 | -35.0409 | -28.3157
18.4654 | -35.0413 | -28.3227
18.4715 | -35.0418 | -28.3298
18.4776 | -35.0422 | -28.3372
18.4837 | -35.0427 | -28.3445
18.4898 | -35.0431 | -28.3517
18.4959 | -35.0436 | -28.3591
18.502 | -35.044 | -28.3664
18.5081 | -35.0445 | -28.3737
18.5142 | -35.0449 | -28.3812
18.5203 | -35.0454 | -28.3884
18.5264 | -35.0458 | -28.3956
18.5325 | -35.0463 | -28.403
18.5386 | -35.0467 | -28.4101
18.5447 | -35.0472 | -28.4173
18.5508 | -35.0477 | -28.4245
18.557 | -35.0481 | -28.4319
18.5631 | -35.0486 | -28.4391
18.5692 | -35.049 | -28.4462
18.5753 | -35.0495 | -28.4533
18.5814 | -35.0499 | -28.4606
18.5875 | -35.0504 | -28.4676
18.5936 | -35.0509 | -28.4748
18.5997 | -35.0513 | -28.4821
18.6058 | -35.0517 | -28.4889
18.6119 | -35.0522 | -28.4961
18.618 | -35.0526 | -28.5032
18.6241 | -35.0531 | -28.5104
18.6302 | -35.0535 | -28.5177
18.6363 | -35.054 | -28.5248
18.6424 | -35.0545 | -28.5316
18.6485 | -35.0549 | -28.5387
18.6546 | -35.0554 | -28.5459
18.6607 | -35.0558 | -28.5532
18.6668 | -35.0563 | -28.5604
18.6729 | -35.0567 | -28.5677
18.679 | -35.0572 | -28.5745
18.6851 | -35.0576 | -28.5815
18.6912 | -35.0581 | -28.5885
18.6974 | -35.0585 | -28.5956
18.7035 | -35.059 | -28.6028
18.7096 | -35.0595 | -28.6102
18.7157 | -35.06 | -28.6175
18.7218 | -35.0604 | -28.6247
18.7279 | -35.0609 | -28.6318
18.734 | -35.0613 | -28.639
18.7401 | -35.0618 | -28.646
18.7462 | -35.0622 | -28.6529
18.7523 | -35.0627 | -28.6602
18.7584 | -35.0632 | -28.6675
18.7645 | -35.0636 | -28.6747
18.7706 | -35.0641 | -28.6816
18.7767 | -35.0646 | -28.6888
18.7828 | -35.065 | -28.6958
18.7889 | -35.0655 | -28.703
18.795 | -35.0659 | -28.7101
18.8011 | -35.0664 | -28.7172
18.8072 | -35.0668 | -28.7246
18.8133 | -35.0673 | -28.7318
18.8194 | -35.0677 | -28.7389
18.8255 | -35.0682 | -28.7461
18.8316 | -35.0687 | -28.7534
18.8377 | -35.0691 | -28.7605
18.8439 | -35.0696 | -28.7676
18.85 | -35.07 | -28.7748
18.8561 | -35.0705 | -28.7817
18.8622 | -35.0709 | -28.7889
18.8683 | -35.0713 | -28.7959
18.8744 | -35.0718 | -28.8031
18.8805 | -35.0722 | -28.8106
18.8866 | -35.0727 | -28.8179
18.8927 | -35.0732 | -28.8249
18.8988 | -35.0736 | -28.8321
18.9049 | -35.0741 | -28.8392
18.911 | -35.0745 | -28.8464
18.9171 | -35.075 | -28.8533
18.9232 | -35.0754 | -28.8606
18.9293 | -35.0759 | -28.8677
18.9354 | -35.0764 | -28.8747
18.9415 | -35.0768 | -28.8818
18.9476 | -35.0773 | -28.8888
18.9537 | -35.0778 | -28.8959
18.9598 | -35.0782 | -28.903
18.9659 | -35.0787 | -28.9102
18.972 | -35.0791 | -28.9176
18.9781 | -35.0796 | -28.9246
18.9843 | -35.08 | -28.9318
18.9904 | -35.0805 | -28.9389
18.9965 | -35.081 | -28.9459
19.0026 | -35.0814 | -28.9529
19.0087 | -35.0819 | -28.9602
19.0148 | -35.0823 | -28.9672
19.0209 | -35.0828 | -28.9746
19.027 | -35.0833 | -28.9817
19.0331 | -35.0837 | -28.9886
19.0392 | -35.0841 | -28.9958
19.0453 | -35.0846 | -29.0028
19.0514 | -35.0851 | -29.0098
19.0575 | -35.0855 | -29.0171
19.0636 | -35.086 | -29.0238
19.0697 | -35.0865 | -29.0306
19.0758 | -35.0869 | -29.0376
19.0819 | -35.0874 | -29.0448
19.088 | -35.0878 | -29.0516
19.0941 | -35.0883 | -29.059
19.1002 | -35.0887 | -29.0659
19.1063 | -35.0891 | -29.073
19.1124 | -35.0896 | -29.08
19.1185 | -35.09 | -29.0872
19.1246 | -35.0905 | -29.0942
19.1308 | -35.0909 | -29.1014
19.1369 | -35.0914 | -29.1084
19.143 | -35.0919 | -29.1153
19.1491 | -35.0923 | -29.1224
19.1552 | -35.0927 | -29.1298
19.1613 | -35.0932 | -29.1369
19.1674 | -35.0936 | -29.144
19.1735 | -35.0941 | -29.1512
19.1796 | -35.0945 | -29.1582
19.1857 | -35.0949 | -29.1652
19.1918 | -35.0954 | -29.1723
19.1979 | -35.0959 | -29.1794
19.204 | -35.0963 | -29.1863
19.2101 | -35.0968 | -29.1932
19.2162 | -35.0972 | -29.2001
19.2223 | -35.0977 | -29.2071
19.2284 | -35.0981 | -29.2141
19.2345 | -35.0986 | -29.2211
19.2406 | -35.099 | -29.2279
19.2467 | -35.0995 | -29.2352
19.2528 | -35.0999 | -29.2426
19.2589 | -35.1004 | -29.2495
19.265 | -35.1009 | -29.2565
19.2712 | -35.1014 | -29.2631
19.2773 | -35.1019 | -29.2701
19.2834 | -35.1023 | -29.277
19.2895 | -35.1028 | -29.2843
19.2956 | -35.1032 | -29.2914
19.3017 | -35.1037 | -29.2984
19.3078 | -35.1041 | -29.3054
19.3139 | -35.1046 | -29.3122
19.32 | -35.105 | -29.3191
19.3261 | -35.1055 | -29.3261
19.3322 | -35.1059 | -29.3329
19.3383 | -35.1063 | -29.3399
19.3444 | -35.1068 | -29.3469
19.3505 | -35.1073 | -29.354
19.3566 | -35.1077 | -29.361
19.3627 | -35.1082 | -29.3681
19.3688 | -35.1086 | -29.3751
19.3749 | -35.1091 | -29.3819
19.381 | -35.1095 | -29.3885
19.3871 | -35.11 | -29.3955
19.3932 | -35.1104 | -29.4026
19.3993 | -35.1108 | -29.4097
19.4054 | -35.1113 | -29.4168
19.4115 | -35.1117 | -29.4238
19.4177 | -35.1121 | -29.4305
19.4238 | -35.1126 | -29.4377
19.4299 | -35.113 | -29.4445
19.436 | -35.1134 | -29.4518
19.4421 | -35.1139 | -29.4588
19.4482 | -35.1143 | -29.4658
19.4543 | -35.1148 | -29.4724
19.4604 | -35.1152 | -29.4795
19.4665 | -35.1156 | -29.4864
19.4726 | -35.116 | -29.4934
19.4787 | -35.1165 | -29.5004
19.4848 | -35.1169 | -29.5075
19.4909 | -35.1174 | -29.5144
19.497 | -35.1178 | -29.5214
19.5031 | -35.1183 | -29.5281
19.5092 | -35.1187 | -29.5349
19.5153 | -35.1192 | -29.5418
19.5214 | -35.1197 | -29.5484
19.5275 | -35.1202 | -29.5555
19.5336 | -35.1206 | -29.5623
19.5397 | -35.121 | -29.5692
19.5458 | -35.1215 | -29.5761
19.5519 | -35.1219 | -29.5831
19.5581 | -35.1223 | -29.5895
19.5642 | -35.1228 | -29.5965
19.5703 | -35.1232 | -29.6034
19.5764 | -35.1236 | -29.6104
19.5825 | -35.1241 | -29.6172
19.5886 | -35.1246 | -29.6242
19.5947 | -35.125 | -29.6309
19.6008 | -35.1255 | -29.6377
19.6069 | -35.1259 | -29.6449
19.613 | -35.1264 | -29.652
19.6191 | -35.1268 | -29.6589
19.6252 | -35.1272 | -29.6657
19.6313 | -35.1276 | -29.6723
19.6374 | -35.1281 | -29.6793
19.6435 | -35.1285 | -29.6864
19.6496 | -35.1289 | -29.6934
19.6557 | -35.1294 | -29.7003
19.6618 | -35.1298 | -29.7068
19.6679 | -35.1302 | -29.7136
19.674 | -35.1307 | -29.7205
19.6801 | -35.1311 | -29.7275
19.6862 | -35.1316 | -29.7349
19.6923 | -35.132 | -29.7416
19.6984 | -35.1324 | -29.7483
19.7046 | -35.1329 | -29.7555
19.7107 | -35.1333 | -29.7622
19.7168 | -35.1337 | -29.7691
19.7229 | -35.1342 | -29.7761
19.729 | -35.1346 | -29.7829
19.7351 | -35.135 | -29.7894
19.7412 | -35.1355 | -29.7965
19.7473 | -35.1359 | -29.8035
19.7534 | -35.1364 | -29.8105
19.7595 | -35.1368 | -29.8175
19.7656 | -35.1372 | -29.8244
19.7717 | -35.1376 | -29.8309
19.7778 | -35.1381 | -29.8379
19.7839 | -35.1386 | -29.8447
19.79 | -35.139 | -29.8516
19.7961 | -35.1394 | -29.8585
19.8022 | -35.1398 | -29.8651
19.8083 | -35.1403 | -29.8721
19.8144 | -35.1408 | -29.8789
19.8205 | -35.1412 | -29.8858
19.8266 | -35.1417 | -29.8925
19.8327 | -35.1421 | -29.8994
19.8388 | -35.1425 | -29.9061
19.845 | -35.143 | -29.913
19.8511 | -35.1434 | -29.9197
19.8572 | -35.1438 | -29.9264
19.8633 | -35.1442 | -29.9333
19.8694 | -35.1446 | -29.9403
19.8755 | -35.1451 | -29.9468
19.8816 | -35.1455 | -29.9539
19.8877 | -35.1459 | -29.9608
19.8938 | -35.1464 | -29.9677
19.8999 | -35.1468 | -29.9747
19.906 | -35.1472 | -29.9815
19.9121 | -35.1477 | -29.9881
19.9182 | -35.1481 | -29.995
19.9243 | -35.1486 | -30.0017
19.9304 | -35.149 | -30.0085
19.9365 | -35.1494 | -30.0154
19.9426 | -35.1499 | -30.022
19.9487 | -35.1503 | -30.0287
19.9548 | -35.1508 | -30.0354
19.9609 | -35.1512 | -30.0422
19.967 | -35.1516 | -30.0488
19.9731 | -35.1521 | -30.0556
19.9792 | -35.1525 | -30.062
19.9853 | -35.153 | -30.0687
19.9915 | -35.1534 | -30.0755
19.9976 | -35.1538 | -30.0822
20.0037 | -35.1542 | -30.089
20.0098 | -35.1547 | -30.0957
20.0159 | -35.1551 | -30.1023
20.022 | -35.1555 | -30.1091
20.0281 | -35.1559 | -30.1159
20.0342 | -35.1564 | -30.1228
20.0403 | -35.1568 | -30.1296
20.0464 | -35.1572 | -30.1364
20.0525 | -35.1577 | -30.143
20.0586 | -35.1581 | -30.1498
20.0647 | -35.1586 | -30.1568
20.0708 | -35.159 | -30.1635
20.0769 | -35.1594 | -30.1702
20.083 | -35.1598 | -30.1769
20.0891 | -35.1602 | -30.1839
20.0952 | -35.1607 | -30.1904
20.1013 | -35.1611 | -30.1972
20.1074 | -35.1616 | -30.204
20.1135 | -35.162 | -30.2108
20.1196 | -35.1625 | -30.2174
20.1257 | -35.1629 | -30.2241
20.1319 | -35.1634 | -30.2309
20.138 | -35.1638 | -30.2377
20.1441 | -35.1643 | -30.2444
20.1502 | -35.1647 | -30.2512
20.1563 | -35.1652 | -30.2578
20.1624 | -35.1656 | -30.2645
20.1685 | -35.166 | -30.2712
20.1746 | -35.1664 | -30.2781
20.1807 | -35.1669 | -30.2849
20.1868 | -35.1673 | -30.2919
20.1929 | -35.1677 | -30.298
20.199 | -35.1682 | -30.3047
20.2051 | -35.1686 | -30.3113
20.2112 | -35.1691 | -30.3181
20.2173 | -35.1695 | -30.3248
20.2234 | -35.17 | -30.3313
20.2295 | -35.1705 | -30.3376
20.2356 | -35.171 | -30.3444
20.2417 | -35.1714 | -30.351
20.2478 | -35.1719 | -30.3577
20.2539 | -35.1723 | -30.3645
20.26 | -35.1727 | -30.371
20.2661 | -35.1732 | -30.3779
20.2723 | -35.1736 | -30.3845
20.2784 | -35.1741 | -30.3911
20.2845 | -35.1745 | -30.3979
20.2906 | -35.175 | -30.4047
20.2967 | -35.1754 | -30.4113
20.3028 | -35.1758 | -30.4179
20.3089 | -35.1763 | -30.4248
20.315 | -35.1767 | -30.4315
20.3211 | -35.1771 | -30.4381
20.3272 | -35.1776 | -30.4449
20.3333 | -35.178 | -30.4514
20.3394 | -35.1784 | -30.4581
20.3455 | -35.1789 | -30.4647
20.3516 | -35.1793 | -30.4713
20.3577 | -35.1797 | -30.4779
20.3638 | -35.1801 | -30.4847
20.3699 | -35.1806 | -30.4908
20.376 | -35.181 | -30.4974
20.3821 | -35.1814 | -30.5042
20.3882 | -35.1819 | -30.5109
20.3943 | -35.1823 | -30.5177
20.4004 | -35.1827 | -30.524
20.4065 | -35.1832 | -30.5305
20.4126 | -35.1836 | -30.5371
20.4188 | -35.184 | -30.544
20.4249 | -35.1845 | -30.5504
20.431 | -35.1849 | -30.557
20.4371 | -35.1853 | -30.5635
20.4432 | -35.1858 | -30.5701
20.4493 | -35.1862 | -30.5769
20.4554 | -35.1867 | -30.5836
20.4615 | -35.1871 | -30.5902
20.4676 | -35.1875 | -30.5967
20.4737 | -35.1879 | -30.6033
20.4798 | -35.1884 | -30.6099
20.4859 | -35.1888 | -30.6167
20.492 | -35.1893 | -30.6231
20.4981 | -35.1897 | -30.6298
20.5042 | -35.1901 | -30.6362
20.5103 | -35.1905 | -30.6427
20.5164 | -35.191 | -30.6493
20.5225 | -35.1914 | -30.6557
20.5286 | -35.1918 | -30.6623
20.5347 | -35.1923 | -30.6685
20.5408 | -35.1927 | -30.6747
20.5469 | -35.1932 | -30.6812
20.553 | -35.1936 | -30.6879
20.5592 | -35.194 | -30.6945
20.5653 | -35.1945 | -30.7012
20.5714 | -35.1949 | -30.7079
20.5775 | -35.1954 | -30.714
20.5836 | -35.1958 | -30.7207
20.5897 | -35.1962 | -30.7274
20.5958 | -35.1967 | -30.7341
20.6019 | -35.1971 | -30.7409
20.608 | -35.1975 | -30.7475
20.6141 | -35.198 | -30.7541
20.6202 | -35.1984 | -30.7605
20.6263 | -35.1989 | -30.767
20.6324 | -35.1993 | -30.7736
20.6385 | -35.1998 | -30.7801
20.6446 | -35.2002 | -30.7865
20.6507 | -35.2006 | -30.7929
20.6568 | -35.2011 | -30.7994
20.6629 | -35.2015 | -30.8063
20.669 | -35.2019 | -30.8127
20.6751 | -35.2024 | -30.8191
20.6812 | -35.2028 | -30.8254
20.6873 | -35.2032 | -30.8321
20.6934 | -35.2036 | -30.8386
20.6995 | -35.2041 | -30.8452
20.7057 | -35.2045 | -30.8517
20.7118 | -35.205 | -30.8584
20.7179 | -35.2055 | -30.8648
20.724 | -35.2059 | -30.8718
20.7301 | -35.2063 | -30.8784
20.7362 | -35.2067 | -30.8849
20.7423 | -35.2072 | -30.8913
20.7484 | -35.2076 | -30.8978
20.7545 | -35.2081 | -30.904
20.7606 | -35.2085 | -30.9109
20.7667 | -35.209 | -30.9174
20.7728 | -35.2094 | -30.924
20.7789 | -35.2098 | -30.9306
20.785 | -35.2102 | -30.937
20.7911 | -35.2107 | -30.9433
20.7972 | -35.2111 | -30.95
20.8033 | -35.2116 | -30.9567
20.8094 | -35.212 | -30.9631
20.8155 | -35.2125 | -30.9697
20.8216 | -35.2129 | -30.9762
20.8277 | -35.2133 | -30.9828
20.8338 | -35.2137 | -30.9891
20.8399 | -35.2142 | -30.9955
20.8461 | -35.2146 | -31.0024
20.8522 | -35.2151 | -31.009
20.8583 | -35.2155 | -31.0154
20.8644 | -35.216 | -31.0219
20.8705 | -35.2164 | -31.0285
20.8766 | -35.2168 | -31.0351
20.8827 | -35.2173 | -31.0417
20.8888 | -35.2177 | -31.0481
20.8949 | -35.2182 | -31.0542
20.901 | -35.2186 | -31.0607
20.9071 | -35.219 | -31.0671
20.9132 | -35.2195 | -31.0737
20.9193 | -35.2199 | -31.0804
20.9254 | -35.2203 | -31.0871
20.9315 | -35.2208 | -31.0933
20.9376 | -35.2212 | -31.0997
20.9437 | -35.2216 | -31.1062
20.9498 | -35.222 | -31.1125
20.9559 | -35.2224 | -31.1189
20.962 | -35.2229 | -31.1252
20.9681 | -35.2233 | -31.1317
20.9742 | -35.2237 | -31.1382
20.9803 | -35.2241 | -31.1445
20.9864 | -35.2246 | -31.151
20.9926 | -35.225 | -31.1575
20.9987 | -35.2254 | -31.1636
21.0048 | -35.2259 | -31.1699
21.0109 | -35.2263 | -31.1764
21.017 | -35.2268 | -31.1827
21.0231 | -35.2272 | -31.1891
21.0292 | -35.2277 | -31.1957
21.0353 | -35.2281 | -31.2021
21.0414 | -35.2285 | -31.2087
21.0475 | -35.2289 | -31.2152
21.0536 | -35.2293 | -31.2215
21.0597 | -35.2298 | -31.2275
21.0658 | -35.2302 | -31.234
21.0719 | -35.2307 | -31.2404
21.078 | -35.2311 | -31.2471
21.0841 | -35.2315 | -31.2534
21.0902 | -35.2319 | -31.2597
21.0963 | -35.2324 | -31.2658
21.1024 | -35.2327 | -31.2719
21.1085 | -35.2332 | -31.2779
21.1146 | -35.2337 | -31.2845
21.1207 | -35.2341 | -31.2907
21.1268 | -35.2346 | -31.2973
21.133 | -35.235 | -31.3037
21.1391 | -35.2354 | -31.3101
21.1452 | -35.2359 | -31.3165
21.1513 | -35.2363 | -31.3228
21.1574 | -35.2367 | -31.3292
21.1635 | -35.2372 | -31.3357
21.1696 | -35.2376 | -31.3419
21.1757 | -35.238 | -31.348
21.1818 | -35.2384 | -31.3542
21.1879 | -35.2388 | -31.3605
21.194 | -35.2393 | -31.3667
21.2001 | -35.2397 | -31.3729
21.2062 | -35.2401 | -31.3793
21.2123 | -35.2406 | -31.3856
21.2184 | -35.241 | -31.3918
21.2245 | -35.2414 | -31.398
21.2306 | -35.2419 | -31.4042
21.2367 | -35.2423 | -31.4108
21.2428 | -35.2427 | -31.417
21.2489 | -35.2431 | -31.4229
21.255 | -35.2436 | -31.4294
21.2611 | -35.244 | -31.4358
21.2672 | -35.2445 | -31.4422
21.2733 | -35.2449 | -31.4486
21.2795 | -35.2453 | -31.4546
21.2856 | -35.2458 | -31.4605
21.2917 | -35.2462 | -31.467
21.2978 | -35.2466 | -31.4733
21.3039 | -35.247 | -31.4794
21.31 | -35.2475 | -31.4863
21.3161 | -35.2478 | -31.4924
21.3222 | -35.2483 | -31.4988
21.3283 | -35.2487 | -31.5051
21.3344 | -35.2491 | -31.5112
21.3405 | -35.2495 | -31.5173
21.3466 | -35.25 | -31.5237
21.3527 | -35.2505 | -31.5296
21.3588 | -35.2509 | -31.5359
21.3649 | -35.2513 | -31.542
21.371 | -35.2518 | -31.5483
21.3771 | -35.2522 | -31.5544
21.3832 | -35.2527 | -31.5608
21.3893 | -35.2531 | -31.5669
21.3954 | -35.2535 | -31.5734
21.4015 | -35.254 | -31.5796
21.4076 | -35.2544 | -31.5861
21.4137 | -35.2549 | -31.5924
21.4199 | -35.2554 | -31.599
21.426 | -35.2558 | -31.6047
21.4321 | -35.2562 | -31.6111
21.4382 | -35.2566 | -31.6177
21.4443 | -35.257 | -31.6239
21.4504 | -35.2574 | -31.6299
21.4565 | -35.2578 | -31.6358
21.4626 | -35.2582 | -31.6421
21.4687 | -35.2587 | -31.6484
21.4748 | -35.2591 | -31.6547
21.4809 | -35.2595 | -31.661
21.487 | -35.2599 | -31.6675
21.4931 | -35.2604 | -31.6735
21.4992 | -35.2608 | -31.6798
21.5053 | -35.2612 | -31.686
21.5114 | -35.2616 | -31.6922
21.5175 | -35.262 | -31.6985
21.5236 | -35.2624 | -31.7047
21.5297 | -35.2629 | -31.7105
21.5358 | -35.2633 | -31.7168
21.5419 | -35.2637 | -31.7228
21.548 | -35.2641 | -31.7289
21.5541 | -35.2645 | -31.7351
21.5602 | -35.265 | -31.7412
21.5664 | -35.2654 | -31.7475
21.5725 | -35.2659 | -31.7538
21.5786 | -35.2663 | -31.7603
21.5847 | -35.2667 | -31.7668
21.5908 | -35.2671 | -31.7728
21.5969 | -35.2675 | -31.7789
21.603 | -35.268 | -31.7852
21.6091 | -35.2684 | -31.7915
21.6152 | -35.2688 | -31.7979
21.6213 | -35.2693 | -31.8042
21.6274 | -35.2697 | -31.8103
21.6335 | -35.2702 | -31.8162
21.6396 | -35.2706 | -31.8226
21.6457 | -35.271 | -31.8287
21.6518 | -35.2715 | -31.8351
21.6579 | -35.2719 | -31.841
21.664 | -35.2724 | -31.8471
21.6701 | -35.2728 | -31.8532
21.6762 | -35.2732 | -31.8594
21.6823 | -35.2736 | -31.8657
21.6884 | -35.274 | -31.8718
21.6945 | -35.2744 | -31.8778
21.7006 | -35.2749 | -31.884
21.7068 | -35.2753 | -31.8898
21.7129 | -35.2758 | -31.8959
21.719 | -35.2763 | -31.9021
21.7251 | -35.2767 | -31.908
21.7312 | -35.2771 | -31.9142
21.7373 | -35.2776 | -31.9203
21.7434 | -35.278 | -31.9264
21.7495 | -35.2785 | -31.9328
21.7556 | -35.2789 | -31.9391
21.7617 | -35.2793 | -31.9452
21.7678 | -35.2797 | -31.9513
21.7739 | -35.2802 | -31.957
21.78 | -35.2806 | -31.963
21.7861 | -35.2811 | -31.9694
21.7922 | -35.2815 | -31.9756
21.7983 | -35.2819 | -31.9818
21.8044 | -35.2823 | -31.988
21.8105 | -35.2827 | -31.9937
21.8166 | -35.2832 | -31.9997
21.8227 | -35.2835 | -32.0061
21.8288 | -35.284 | -32.0123
21.8349 | -35.2844 | -32.0188
21.841 | -35.2848 | -32.0251
21.8471 | -35.2853 | -32.0311
21.8533 | -35.2857 | -32.0375
21.8594 | -35.2861 | -32.0437
21.8655 | -35.2865 | -32.0501
21.8716 | -35.287 | -32.0561
21.8777 | -35.2874 | -32.0622
21.8838 | -35.2878 | -32.0684
21.8899 | -35.2882 | -32.0746
21.896 | -35.2887 | -32.0809
21.9021 | -35.2891 | -32.0871
21.9082 | -35.2895 | -32.0935
21.9143 | -35.29 | -32.0995
21.9204 | -35.2904 | -32.1056
21.9265 | -35.2908 | -32.1118
21.9326 | -35.2913 | -32.118
21.9387 | -35.2917 | -32.1243
21.9448 | -35.2921 | -32.1306
21.9509 | -35.2926 | -32.1368
21.957 | -35.293 | -32.143
21.9631 | -35.2934 | -32.1491
21.9692 | -35.2938 | -32.1552
21.9753 | -35.2942 | -32.1616
21.9814 | -35.2947 | -32.1678
21.9875 | -35.2951 | -32.174
21.9937 | -35.2955 | -32.1802
21.9998 | -35.296 | -32.1865
22.0059 | -35.2964 | -32.1927
22.012 | -35.2968 | -32.1989
22.0181 | -35.2973 | -32.2047
22.0242 | -35.2977 | -32.2109
22.0303 | -35.2981 | -32.2172
22.0364 | -35.2985 | -32.2233
22.0425 | -35.299 | -32.2292
22.0486 | -35.2994 | -32.2353
22.0547 | -35.2998 | -32.2415
22.0608 | -35.3002 | -32.2477
22.0669 | -35.3006 | -32.2538
22.073 | -35.3011 | -32.2597
22.0791 | -35.3015 | -32.2658
22.0852 | -35.3019 | -32.2718
22.0913 | -35.3022 | -32.2778
22.0974 | -35.3026 | -32.2839
22.1035 | -35.3031 | -32.2902
22.1096 | -35.3035 | -32.2964
22.1157 | -35.3039 | -32.3027
22.1218 | -35.3044 | -32.3089
22.1279 | -35.3048 | -32.3146
22.134 | -35.3052 | -32.3208
22.1402 | -35.3057 | -32.327
22.1463 | -35.3061 | -32.3331
22.1524 | -35.3065 | -32.3391
22.1585 | -35.3069 | -32.3451
22.1646 | -35.3073 | -32.3512
22.1707 | -35.3077 | -32.3574
22.1768 | -35.3081 | -32.3634
22.1829 | -35.3085 | -32.3695
22.189 | -35.309 | -32.3757
22.1951 | -35.3094 | -32.3816
22.2012 | -35.3098 | -32.3873
22.2073 | -35.3102 | -32.3935
22.2134 | -35.3106 | -32.3996
22.2195 | -35.311 | -32.4056
22.2256 | -35.3115 | -32.4118
22.2317 | -35.3119 | -32.4177
22.2378 | -35.3123 | -32.4239
22.2439 | -35.3127 | -32.4298
22.25 | -35.3131 | -32.4359
22.2561 | -35.3136 | -32.4419
22.2622 | -35.314 | -32.4479
22.2683 | -35.3144 | -32.4536
22.2744 | -35.3148 | -32.4598
22.2806 | -35.3153 | -32.4659
22.2867 | -35.3156 | -32.4721
22.2928 | -35.3161 | -32.4782
22.2989 | -35.3165 | -32.4842
22.305 | -35.317 | -32.4902
22.3111 | -35.3174 | -32.4964
22.3172 | -35.3178 | -32.5022
22.3233 | -35.3183 | -32.5083
22.3294 | -35.3187 | -32.5141
22.3355 | -35.3191 | -32.5201
22.3416 | -35.3195 | -32.5261
22.3477 | -35.3199 | -32.5321
22.3538 | -35.3203 | -32.538
22.3599 | -35.3207 | -32.5442
22.366 | -35.3212 | -32.55
22.3721 | -35.3216 | -32.5557
22.3782 | -35.322 | -32.5619
22.3843 | -35.3224 | -32.5679
22.3904 | -35.3228 | -32.5743
22.3965 | -35.3232 | -32.5802
22.4026 | -35.3237 | -32.5861
22.4087 | -35.3241 | -32.592
22.4148 | -35.3245 | -32.5979
22.4209 | -35.3249 | -32.604
22.4271 | -35.3253 | -32.6102
22.4332 | -35.3257 | -32.616
22.4393 | -35.3261 | -32.622
22.4454 | -35.3266 | -32.6278
22.4515 | -35.327 | -32.6338
22.4576 | -35.3274 | -32.6398
22.4637 | -35.3278 | -32.646
22.4698 | -35.3282 | -32.652
22.4759 | -35.3287 | -32.658
22.482 | -35.329 | -32.6644
22.4881 | -35.3294 | -32.6704
22.4942 | -35.3298 | -32.6765
22.5003 | -35.3303 | -32.6827
22.5064 | -35.3307 | -32.6889
22.5125 | -35.3311 | -32.6948
22.5186 | -35.3315 | -32.7008
22.5247 | -35.332 | -32.7069
22.5308 | -35.3324 | -32.7131
22.5369 | -35.3328 | -32.7192
22.543 | -35.3331 | -32.7252
22.5491 | -35.3335 | -32.7311
22.5552 | -35.334 | -32.7371
22.5613 | -35.3344 | -32.7434
22.5675 | -35.3348 | -32.7493
22.5736 | -35.3352 | -32.7554
22.5797 | -35.3357 | -32.7616
22.5858 | -35.3361 | -32.7674
22.5919 | -35.3365 | -32.7733
22.598 | -35.337 | -32.7796
22.6041 | -35.3374 | -32.7855
22.6102 | -35.3378 | -32.7917
22.6163 | -35.3382 | -32.7978
22.6224 | -35.3387 | -32.8041
22.6285 | -35.3391 | -32.8101
22.6346 | -35.3395 | -32.8161
22.6407 | -35.34 | -32.8221
22.6468 | -35.3404 | -32.8281
22.6529 | -35.3408 | -32.8339
22.659 | -35.3412 | -32.8398
22.6651 | -35.3416 | -32.846
22.6712 | -35.3421 | -32.852
22.6773 | -35.3425 | -32.8579
22.6834 | -35.3429 | -32.864
22.6895 | -35.3433 | -32.87
22.6956 | -35.3437 | -32.8762
22.7017 | -35.3441 | -32.8821
22.7079 | -35.3445 | -32.8878
22.714 | -35.3449 | -32.8939
22.7201 | -35.3454 | -32.8999
22.7262 | -35.3457 | -32.9057
22.7323 | -35.3462 | -32.9116
22.7384 | -35.3466 | -32.9177
22.7445 | -35.347 | -32.9237
22.7506 | -35.3474 | -32.9296
22.7567 | -35.3478 | -32.9356
22.7628 | -35.3482 | -32.9412
22.7689 | -35.3486 | -32.9472
22.775 | -35.349 | -32.9532
22.7811 | -35.3494 | -32.9593
22.7872 | -35.3499 | -32.9654
22.7933 | -35.3503 | -32.9712
22.7994 | -35.3507 | -32.9771
22.8055 | -35.3511 | -32.9833
22.8116 | -35.3515 | -32.989
22.8177 | -35.3519 | -32.9951
22.8238 | -35.3524 | -33.0012
22.8299 | -35.3528 | -33.0072
22.836 | -35.3532 | -33.0133
22.8421 | -35.3536 | -33.0193
22.8482 | -35.354 | -33.0253
22.8544 | -35.3543 | -33.0314
22.8605 | -35.3547 | -33.0374
22.8666 | -35.3551 | -33.0432
22.8727 | -35.3555 | -33.0491
22.8788 | -35.356 | -33.0548
22.8849 | -35.3564 | -33.0608
22.891 | -35.3568 | -33.0669
22.8971 | -35.3572 | -33.0729
22.9032 | -35.3576 | -33.0787
22.9093 | -35.358 | -33.0847
22.9154 | -35.3584 | -33.0907
22.9215 | -35.3588 | -33.0967
22.9276 | -35.3592 | -33.1027
22.9337 | -35.3596 | -33.1085
22.9398 | -35.36 | -33.1144
22.9459 | -35.3604 | -33.1208
22.952 | -35.3609 | -33.1267
22.9581 | -35.3612 | -33.1328
22.9642 | -35.3617 | -33.139
22.9703 | -35.3621 | -33.1449
22.9764 | -35.3625 | -33.1508
22.9825 | -35.3629 | -33.1568
22.9886 | -35.3632 | -33.163
22.9948 | -35.3636 | -33.169
23.0009 | -35.364 | -33.1752
23.007 | -35.3645 | -33.1811
23.0131 | -35.3649 | -33.1871
23.0192 | -35.3653 | -33.193
23.0253 | -35.3657 | -33.199
23.0314 | -35.3661 | -33.205
23.0375 | -35.3665 | -33.211
23.0436 | -35.3669 | -33.2166
23.0497 | -35.3673 | -33.2227
23.0558 | -35.3677 | -33.2287
23.0619 | -35.3681 | -33.2346
23.068 | -35.3685 | -33.2406
23.0741 | -35.3689 | -33.2461
23.0802 | -35.3693 | -33.2521
23.0863 | -35.3697 | -33.2581
23.0924 | -35.3701 | -33.2642
23.0985 | -35.3705 | -33.2701
23.1046 | -35.3709 | -33.2763
23.1107 | -35.3713 | -33.2822
23.1168 | -35.3717 | -33.2884
23.1229 | -35.3721 | -33.2941
23.129 | -35.3726 | -33.3
23.1351 | -35.3729 | -33.3057
23.1413 | -35.3733 | -33.3116
23.1474 | -35.3737 | -33.3173
23.1535 | -35.3741 | -33.3234
23.1596 | -35.3745 | -33.3294
23.1657 | -35.3749 | -33.335
23.1718 | -35.3753 | -33.3409
23.1779 | -35.3758 | -33.3469
23.184 | -35.3762 | -33.3525
23.1901 | -35.3766 | -33.3586
23.1962 | -35.377 | -33.3646
23.2023 | -35.3775 | -33.3706
23.2084 | -35.3778 | -33.3766
23.2145 | -35.3782 | -33.3822
23.2206 | -35.3786 | -33.3882
23.2267 | -35.379 | -33.3943
23.2328 | -35.3795 | -33.4002
23.2389 | -35.3799 | -33.4062
23.245 | -35.3803 | -33.4122
23.2511 | -35.3807 | -33.4179
23.2572 | -35.3811 | -33.4239
23.2633 | -35.3816 | -33.4299
23.2694 | -35.382 | -33.436
23.2755 | -35.3823 | -33.442
23.2817 | -35.3827 | -33.448
23.2878 | -35.3831 | -33.4539
23.2939 | -35.3836 | -33.4599
23.3 | -35.384 | -33.4659
23.3061 | -35.3844 | -33.4719
23.3122 | -35.3848 | -33.4783
23.3183 | -35.3852 | -33.4844
23.3244 | -35.3856 | -33.4904
23.3305 | -35.386 | -33.4962
23.3366 | -35.3864 | -33.5023
23.3427 | -35.3869 | -33.5082
23.3488 | -35.3873 | -33.5139
23.3549 | -35.3877 | -33.5198
23.361 | -35.3881 | -33.5256
23.3671 | -35.3885 | -33.5315
23.3732 | -35.3889 | -33.5375
23.3793 | -35.3894 | -33.5431
23.3854 | -35.3898 | -33.549
23.3915 | -35.3902 | -33.5548
23.3976 | -35.3906 | -33.5608
23.4037 | -35.391 | -33.5667
23.4098 | -35.3914 | -33.5725
23.4159 | -35.3918 | -33.5785
23.422 | -35.3922 | -33.5846
23.4282 | -35.3926 | -33.5906
23.4343 | -35.393 | -33.5964
23.4404 | -35.3933 | -33.6022
23.4465 | -35.3937 | -33.6082
23.4526 | -35.3941 | -33.6141
23.4587 | -35.3945 | -33.6201
23.4648 | -35.3949 | -33.626
23.4709 | -35.3952 | -33.6319
23.477 | -35.3956 | -33.638
23.4831 | -35.396 | -33.6438
23.4892 | -35.3964 | -33.6497
23.4953 | -35.3969 | -33.6557
23.5014 | -35.3972 | -33.6615
23.5075 | -35.3977 | -33.6674
23.5136 | -35.3981 | -33.6732
23.5197 | -35.3985 | -33.6791
23.5258 | -35.3988 | -33.685
23.5319 | -35.3992 | -33.6906
23.538 | -35.3996 | -33.6965
23.5441 | -35.4 | -33.7025
23.5502 | -35.4003 | -33.7084
23.5563 | -35.4007 | -33.7142
23.5624 | -35.4011 | -33.72
23.5686 | -35.4015 | -33.7258
23.5747 | -35.4018 | -33.7316
23.5808 | -35.4022 | -33.7372
23.5869 | -35.4026 | -33.743
23.593 | -35.403 | -33.7487
23.5991 | -35.4035 | -33.7549
23.6052 | -35.4038 | -33.7606
23.6113 | -35.4042 | -33.7665
23.6174 | -35.4046 | -33.7725
23.6235 | -35.405 | -33.7784
23.6296 | -35.4054 | -33.7843
23.6357 | -35.4057 | -33.7901
23.6418 | -35.4061 | -33.7962
23.6479 | -35.4065 | -33.8021
23.654 | -35.4069 | -33.8081
23.6601 | -35.4073 | -33.8141
23.6662 | -35.4077 | -33.8199
23.6723 | -35.4081 | -33.8257
23.6784 | -35.4085 | -33.8319
23.6845 | -35.4089 | -33.8378
23.6906 | -35.4093 | -33.8438
23.6967 | -35.4096 | -33.8496
23.7028 | -35.41 | -33.8555
23.7089 | -35.4104 | -33.8611
23.7151 | -35.4107 | -33.8669
23.7212 | -35.4111 | -33.8725
23.7273 | -35.4115 | -33.8784
23.7334 | -35.4119 | -33.8844
23.7395 | -35.4123 | -33.8903
23.7456 | -35.4126 | -33.8962
23.7517 | -35.413 | -33.9023
23.7578 | -35.4134 | -33.9076
23.7639 | -35.4138 | -33.9134
23.77 | -35.4141 | -33.9192
23.7761 | -35.4145 | -33.9252
23.7822 | -35.4148 | -33.9308
23.7883 | -35.4152 | -33.9367
23.7944 | -35.4156 | -33.9425
23.8005 | -35.4159 | -33.9484
23.8066 | -35.4163 | -33.9542
23.8127 | -35.4167 | -33.9603
23.8188 | -35.4171 | -33.9663
23.8249 | -35.4175 | -33.9724
23.831 | -35.4178 | -33.9782
23.8371 | -35.4182 | -33.9841
23.8432 | -35.4186 | -33.99
23.8493 | -35.4189 | -33.9964
23.8555 | -35.4193 | -34.0023
23.8616 | -35.4197 | -34.0084
23.8677 | -35.42 | -34.0145
23.8738 | -35.4205 | -34.0204
23.8799 | -35.4209 | -34.0263
23.886 | -35.4212 | -34.0319
23.8921 | -35.4215 | -34.0378
23.8982 | -35.4219 | -34.0436
23.9043 | -35.4223 | -34.0496
23.9104 | -35.4227 | -34.0556
23.9165 | -35.423 | -34.0615
23.9226 | -35.4234 | -34.0672
23.9287 | -35.4237 | -34.0733
23.9348 | -35.4241 | -34.0791
23.9409 | -35.4245 | -34.085
23.947 | -35.4249 | -34.0911
23.9531 | -35.4253 | -34.0968
23.9592 | -35.4256 | -34.1028
23.9653 | -35.426 | -34.1088
23.9714 | -35.4263 | -34.1146
23.9775 | -35.4267 | -34.1205
23.9836 | -35.4271 | -34.1264
23.9897 | -35.4275 | -34.1325
23.9958 | -35.4279 | -34.1382
24.002 | -35.4282 | -34.1442
24.0081 | -35.4286 | -34.1501
24.0142 | -35.429 | -34.1561
24.0203 | -35.4294 | -34.1621
24.0264 | -35.4297 | -34.1678
24.0325 | -35.4301 | -34.1738
24.0386 | -35.4305 | -34.1799
24.0447 | -35.4308 | -34.1857
24.0508 | -35.4312 | -34.1917
24.0569 | -35.4315 | -34.1976
24.063 | -35.432 | -34.2035
24.0691 | -35.4323 | -34.2096
24.0752 | -35.4327 | -34.2152
24.0813 | -35.4332 | -34.2211
24.0874 | -35.4335 | -34.2269
24.0935 | -35.4339 | -34.2331
24.0996 | -35.4342 | -34.2391
24.1057 | -35.4346 | -34.2449
24.1118 | -35.435 | -34.251
24.1179 | -35.4354 | -34.2571
24.124 | -35.4358 | -34.2629
24.1301 | -35.4361 | -34.2687
24.1362 | -35.4365 | -34.2745
24.1424 | -35.4369 | -34.2803
24.1485 | -35.4373 | -34.286
24.1546 | -35.4376 | -34.2915
24.1607 | -35.4379 | -34.2976
24.1668 | -35.4383 | -34.3031
24.1729 | -35.4386 | -34.3089
24.179 | -35.439 | -34.3149
24.1851 | -35.4394 | -34.3208
24.1912 | -35.4398 | -34.3267
24.1973 | -35.4401 | -34.3327
24.2034 | -35.4405 | -34.3382
24.2095 | -35.4409 | -34.344
24.2156 | -35.4412 | -34.35
24.2217 | -35.4416 | -34.3558
24.2278 | -35.442 | -34.3615
24.2339 | -35.4424 | -34.3674
24.24 | -35.4427 | -34.3728
24.2461 | -35.4431 | -34.3785
24.2522 | -35.4434 | -34.3845
24.2583 | -35.4438 | -34.3904
24.2644 | -35.4442 | -34.3963
24.2705 | -35.4446 | -34.4022
24.2766 | -35.445 | -34.4079
24.2827 | -35.4453 | -34.4136
24.2889 | -35.4456 | -34.4196
24.295 | -35.446 | -34.4254
24.3011 | -35.4464 | -34.431
24.3072 | -35.4468 | -34.4371
24.3133 | -35.4471 | -34.4431
24.3194 | -35.4475 | -34.4492
24.3255 | -35.4479 | -34.4551
24.3316 | -35.4482 | -34.461
24.3377 | -35.4485 | -34.4671
24.3438 | -35.4489 | -34.4732
24.3499 | -35.4492 | -34.4789
24.356 | -35.4495 | -34.4847
24.3621 | -35.4499 | -34.4905
24.3682 | -35.4502 | -34.4965
24.3743 | -35.4506 | -34.5023
24.3804 | -35.4509 | -34.508
24.3865 | -35.4513 | -34.5139
24.3926 | -35.4516 | -34.5198
24.3987 | -35.452 | -34.5258
24.4048 | -35.4523 | -34.5317
24.4109 | -35.4527 | -34.5369
24.417 | -35.453 | -34.5427
24.4231 | -35.4534 | -34.5486
24.4293 | -35.4537 | -34.5543
24.4354 | -35.4541 | -34.5603
24.4415 | -35.4545 | -34.5661
24.4476 | -35.4548 | -34.5715
24.4537 | -35.4552 | -34.5773
24.4598 | -35.4556 | -34.5833
24.4659 | -35.4559 | -34.5893
24.472 | -35.4563 | -34.5952
24.4781 | -35.4567 | -34.6012
24.4842 | -35.457 | -34.6069
24.4903 | -35.4574 | -34.6126
24.4964 | -35.4578 | -34.6185
24.5025 | -35.4581 | -34.6243
24.5086 | -35.4585 | -34.6302
24.5147 | -35.4589 | -34.6359
24.5208 | -35.4592 | -34.6416
24.5269 | -35.4596 | -34.6472
24.533 | -35.46 | -34.653
24.5391 | -35.4603 | -34.6589
24.5452 | -35.4607 | -34.6647
24.5513 | -35.4611 | -34.6703
24.5574 | -35.4615 | -34.6761
24.5635 | -35.4618 | -34.6819
24.5696 | -35.4622 | -34.6877
24.5758 | -35.4625 | -34.6934
24.5819 | -35.4629 | -34.6989
24.588 | -35.4632 | -34.7043
24.5941 | -35.4635 | -34.7104
24.6002 | -35.4639 | -34.716
24.6063 | -35.4643 | -34.722
24.6124 | -35.4646 | -34.7281
24.6185 | -35.465 | -34.734
24.6246 | -35.4653 | -34.7398
24.6307 | -35.4657 | -34.7461
24.6368 | -35.466 | -34.7521
24.6429 | -35.4664 | -34.7578
24.649 | -35.4667 | -34.7637
24.6551 | -35.4671 | -34.7696
24.6612 | -35.4674 | -34.775
24.6673 | -35.4678 | -34.7807
24.6734 | -35.4681 | -34.7865
24.6795 | -35.4684 | -34.7924
24.6856 | -35.4688 | -34.7982
24.6917 | -35.4691 | -34.8041
24.6978 | -35.4695 | -34.8098
24.7039 | -35.4699 | -34.8155
24.71 | -35.4702 | -34.8213
24.7162 | -35.4706 | -34.8272
24.7223 | -35.471 | -34.8331
24.7284 | -35.4713 | -34.8388
24.7345 | -35.4716 | -34.8445
24.7406 | -35.4719 | -34.8503
24.7467 | -35.4723 | -34.8563
24.7528 | -35.4727 | -34.8618
24.7589 | -35.473 | -34.8673
24.765 | -35.4735 | -34.8729
24.7711 | -35.4739 | -34.879
24.7772 | -35.4742 | -34.8849
24.7833 | -35.4745 | -34.8909
24.7894 | -35.4748 | -34.897
24.7955 | -35.4752 | -34.9027
24.8016 | -35.4756 | -34.9086
24.8077 | -35.4759 | -34.9144
24.8138 | -35.4763 | -34.9203
24.8199 | -35.4766 | -34.926
24.826 | -35.477 | -34.9316
24.8321 | -35.4773 | -34.9372
24.8382 | -35.4776 | -34.9429
24.8443 | -35.478 | -34.9489
24.8504 | -35.4783 | -34.9544
24.8565 | -35.4787 | -34.9602
24.8627 | -35.479 | -34.966
24.8688 | -35.4793 | -34.9719
24.8749 | -35.4797 | -34.9777
24.881 | -35.48 | -34.9834
24.8871 | -35.4804 | -34.9894
24.8932 | -35.4807 | -34.9954
24.8993 | -35.481 | -35.0012
24.9054 | -35.4814 | -35.0067
24.9115 | -35.4817 | -35.0123
24.9176 | -35.4821 | -35.0181
24.9237 | -35.4825 | -35.0238
24.9298 | -35.4828 | -35.0296
24.9359 | -35.4831 | -35.0356
24.942 | -35.4835 | -35.0414
24.9481 | -35.4838 | -35.0475
24.9542 | -35.4842 | -35.0533
24.9603 | -35.4846 | -35.0589
24.9664 | -35.4849 | -35.0645
24.9725 | -35.4852 | -35.0703
24.9786 | -35.4855 | -35.0758
24.9847 | -35.4858 | -35.0817
24.9908 | -35.4862 | -35.0877
24.9969 | -35.4865 | -35.0939
25.0031 | -35.4868 | -35.1
25.0092 | -35.4871 | -35.1058
25.0153 | -35.4874 | -35.1115
25.0214 | -35.4878 | -35.1175
25.0275 | -35.4881 | -35.1232
25.0336 | -35.4885 | -35.1289
25.0397 | -35.4888 | -35.135
25.0458 | -35.4892 | -35.1406
25.0519 | -35.4895 | -35.1465
25.058 | -35.4898 | -35.1522
25.0641 | -35.4902 | -35.1579
25.0702 | -35.4905 | -35.1635
25.0763 | -35.4909 | -35.1695
25.0824 | -35.4912 | -35.1754
25.0885 | -35.4916 | -35.181
25.0946 | -35.4919 | -35.1868
25.1007 | -35.4922 | -35.1928
25.1068 | -35.4926 | -35.1984
25.1129 | -35.4929 | -35.2043
25.119 | -35.4933 | -35.2099
25.1251 | -35.4936 | -35.2158
25.1312 | -35.494 | -35.2216
25.1373 | -35.4943 | -35.2273
25.1435 | -35.4947 | -35.233
25.1496 | -35.4951 | -35.2384
25.1557 | -35.4954 | -35.2442
25.1618 | -35.4958 | -35.25
25.1679 | -35.4961 | -35.2561
25.174 | -35.4964 | -35.2617
25.1801 | -35.4967 | -35.2675
25.1862 | -35.4971 | -35.2731
25.1923 | -35.4974 | -35.279
25.1984 | -35.4978 | -35.2849
25.2045 | -35.4981 | -35.2906
25.2106 | -35.4984 | -35.2963
25.2167 | -35.4987 | -35.3018
25.2228 | -35.499 | -35.3075
25.2289 | -35.4993 | -35.3135
25.235 | -35.4996 | -35.3193
25.2411 | -35.5 | -35.3254
25.2472 | -35.5004 | -35.3313
25.2533 | -35.5007 | -35.3367
25.2594 | -35.501 | -35.343
25.2655 | -35.5013 | -35.3492
25.2716 | -35.5017 | -35.3546
25.2777 | -35.502 | -35.3607
25.2838 | -35.5023 | -35.3667
25.29 | -35.5027 | -35.3725
25.2961 | -35.503 | -35.3783
25.3022 | -35.5034 | -35.3843
25.3083 | -35.5037 | -35.3899
25.3144 | -35.504 | -35.3956
25.3205 | -35.5043 | -35.4015
25.3266 | -35.5046 | -35.4071
25.3327 | -35.5049 | -35.413
25.3388 | -35.5052 | -35.4189
25.3449 | -35.5055 | -35.4243
25.351 | -35.5058 | -35.4303
25.3571 | -35.5061 | -35.4362
25.3632 | -35.5064 | -35.4419
25.3693 | -35.5067 | -35.4475
25.3754 | -35.507 | -35.4532
25.3815 | -35.5074 | -35.459
25.3876 | -35.5077 | -35.4644
25.3937 | -35.508 | -35.4703
25.3998 | -35.5083 | -35.4756
25.4059 | -35.5087 | -35.4815
25.412 | -35.509 | -35.4873
25.4181 | -35.5094 | -35.4932
25.4242 | -35.5097 | -35.499
25.4304 | -35.51 | -35.5047
25.4365 | -35.5103 | -35.5106
25.4426 | -35.5106 | -35.5162
25.4487 | -35.511 | -35.522
25.4548 | -35.5113 | -35.528
25.4609 | -35.5116 | -35.5337
25.467 | -35.512 | -35.5392
25.4731 | -35.5123 | -35.5448
25.4792 | -35.5126 | -35.5506
25.4853 | -35.513 | -35.5563
25.4914 | -35.5133 | -35.562
25.4975 | -35.5136 | -35.5677
25.5036 | -35.514 | -35.573
25.5097 | -35.5143 | -35.5785
25.5158 | -35.5147 | -35.5841
25.5219 | -35.515 | -35.5899
25.528 | -35.5154 | -35.5957
25.5341 | -35.5157 | -35.6015
25.5402 | -35.5161 | -35.6067
25.5463 | -35.5164 | -35.6125
25.5524 | -35.5168 | -35.6185
25.5585 | -35.5171 | -35.624
25.5646 | -35.5175 | -35.6301
25.5707 | -35.5177 | -35.6356
25.5769 | -35.518 | -35.6412
25.583 | -35.5184 | -35.647
25.5891 | -35.5187 | -35.6526
25.5952 | -35.5191 | -35.6584
25.6013 | -35.5194 | -35.6643
25.6074 | -35.5197 | -35.6697
25.6135 | -35.52 | -35.6754
25.6196 | -35.5204 | -35.6811
25.6257 | -35.5207 | -35.6866
25.6318 | -35.521 | -35.6925
25.6379 | -35.5213 | -35.698
25.644 | -35.5215 | -35.7036
25.6501 | -35.5219 | -35.7093
25.6562 | -35.5222 | -35.7152
25.6623 | -35.5225 | -35.7211
25.6684 | -35.5229 | -35.7268
25.6745 | -35.5232 | -35.7326
25.6806 | -35.5235 | -35.7378
25.6867 | -35.5239 | -35.7435
25.6928 | -35.5242 | -35.7492
25.6989 | -35.5245 | -35.7551
25.705 | -35.5248 | -35.7609
25.7111 | -35.5251 | -35.7664
25.7173 | -35.5254 | -35.7717
25.7234 | -35.5258 | -35.7774
25.7295 | -35.5261 | -35.783
25.7356 | -35.5265 | -35.7891
25.7417 | -35.5268 | -35.7945
25.7478 | -35.527 | -35.7997
25.7539 | -35.5274 | -35.8055
25.76 | -35.5277 | -35.8113
25.7661 | -35.528 | -35.8169
25.7722 | -35.5283 | -35.8228
25.7783 | -35.5286 | -35.8284
25.7844 | -35.529 | -35.834
25.7905 | -35.5293 | -35.8399
25.7966 | -35.5296 | -35.8452
25.8027 | -35.5299 | -35.8511
25.8088 | -35.5302 | -35.8568
25.8149 | -35.5305 | -35.8623
25.821 | -35.5308 | -35.868
25.8271 | -35.5311 | -35.8737
25.8332 | -35.5315 | -35.8794
25.8393 | -35.5318 | -35.8851
25.8454 | -35.5321 | -35.8907
25.8515 | -35.5324 | -35.8964
25.8576 | -35.5328 | -35.902
25.8638 | -35.5331 | -35.9077
25.8699 | -35.5334 | -35.9132
25.876 | -35.5337 | -35.9188
25.8821 | -35.534 | -35.9243
25.8882 | -35.5344 | -35.93
25.8943 | -35.5347 | -35.9359
25.9004 | -35.535 | -35.9415
25.9065 | -35.5353 | -35.947
25.9126 | -35.5357 | -35.9527
25.9187 | -35.536 | -35.9581
25.9248 | -35.5363 | -35.9637
25.9309 | -35.5366 | -35.9694
25.937 | -35.5369 | -35.975
25.9431 | -35.5372 | -35.9808
25.9492 | -35.5375 | -35.9864
25.9553 | -35.5378 | -35.9922
25.9614 | -35.5382 | -35.9977
25.9675 | -35.5385 | -36.0031
25.9736 | -35.5389 | -36.0089
25.9797 | -35.5391 | -36.0145
25.9858 | -35.5394 | -36.0201
25.9919 | -35.5397 | -36.0259
25.998 | -35.54 | -36.0315
26.0042 | -35.5403 | -36.0374
26.0103 | -35.5405 | -36.0431
26.0164 | -35.5408 | -36.0485
26.0225 | -35.5411 | -36.0543
26.0286 | -35.5415 | -36.0597
26.0347 | -35.5417 | -36.0656
26.0408 | -35.542 | -36.0714
26.0469 | -35.5423 | -36.0764
26.053 | -35.5426 | -36.0816
26.0591 | -35.5429 | -36.0877
26.0652 | -35.5432 | -36.093
26.0713 | -35.5435 | -36.0988
26.0774 | -35.5437 | -36.1044
26.0835 | -35.5441 | -36.1099
26.0896 | -35.5444 | -36.1154
26.0957 | -35.5447 | -36.1209
26.1018 | -35.545 | -36.1264
26.1079 | -35.5453 | -36.1319
26.114 | -35.5456 | -36.1373
26.1201 | -35.5459 | -36.1427
26.1262 | -35.5463 | -36.1482
26.1323 | -35.5466 | -36.1536
26.1384 | -35.5469 | -36.1587
26.1445 | -35.5472 | -36.1645
26.1507 | -35.5475 | -36.1701
26.1568 | -35.5478 | -36.1757
26.1629 | -35.5481 | -36.1815
26.169 | -35.5484 | -36.1873
26.1751 | -35.5487 | -36.1927
26.1812 | -35.549 | -36.1987
26.1873 | -35.5493 | -36.2042
26.1934 | -35.5496 | -36.2095
26.1995 | -35.55 | -36.215
26.2056 | -35.5502 | -36.2205
26.2117 | -35.5505 | -36.226
26.2178 | -35.5508 | -36.2318
26.2239 | -35.5511 | -36.2375
26.23 | -35.5514 | -36.2433
26.2361 | -35.5517 | -36.2491
26.2422 | -35.5519 | -36.2544
26.2483 | -35.5523 | -36.2598
26.2544 | -35.5526 | -36.2651
26.2605 | -35.5529 | -36.2708
26.2666 | -35.5533 | -36.2762
26.2727 | -35.5536 | -36.2819
26.2788 | -35.5538 | -36.2874
26.2849 | -35.554 | -36.293
26.2911 | -35.5543 | -36.2984
26.2972 | -35.5546 | -36.3035
26.3033 | -35.5549 | -36.3089
26.3094 | -35.5552 | -36.3145
26.3155 | -35.5555 | -36.32
26.3216 | -35.5558 | -36.3253
26.3277 | -35.5561 | -36.3309
26.3338 | -35.5564 | -36.3362
26.3399 | -35.5568 | -36.3417
26.346 | -35.557 | -36.3464
26.3521 | -35.5573 | -36.3522
26.3582 | -35.5576 | -36.3576
26.3643 | -35.5579 | -36.363
26.3704 | -35.5582 | -36.3683
26.3765 | -35.5585 | -36.3741
26.3826 | -35.5589 | -36.3791
26.3887 | -35.5592 | -36.3845
26.3948 | -35.5595 | -36.39
26.4009 | -35.5599 | -36.3955
26.407 | -35.5601 | -36.401
26.4131 | -35.5604 | -36.4062
26.4192 | -35.5607 | -36.4114
26.4253 | -35.561 | -36.417
26.4314 | -35.5613 | -36.4225
26.4376 | -35.5616 | -36.4281
26.4437 | -35.5619 | -36.4333
26.4498 | -35.5622 | -36.4388
26.4559 | -35.5625 | -36.4445
26.462 | -35.5628 | -36.4504
26.4681 | -35.5631 | -36.4556
26.4742 | -35.5634 | -36.4613
26.4803 | -35.5637 | -36.4665
26.4864 | -35.5639 | -36.4717
26.4925 | -35.5642 | -36.4768
26.4986 | -35.5645 | -36.4824
26.5047 | -35.5648 | -36.4877
26.5108 | -35.5651 | -36.4934
26.5169 | -35.5654 | -36.4986
26.523 | -35.5658 | -36.5035
26.5291 | -35.566 | -36.5095
26.5352 | -35.5663 | -36.5152
26.5413 | -35.5666 | -36.521
26.5474 | -35.567 | -36.5264
26.5535 | -35.5673 | -36.5321
26.5596 | -35.5676 | -36.537
26.5657 | -35.5678 | -36.5426
26.5718 | -35.5682 | -36.5482
26.578 | -35.5685 | -36.5538
26.5841 | -35.5687 | -36.5592
26.5902 | -35.569 | -36.5647
26.5963 | -35.5693 | -36.5701
26.6024 | -35.5696 | -36.5758
26.6085 | -35.5699 | -36.5813
26.6146 | -35.5701 | -36.5865
26.6207 | -35.5704 | -36.5917
26.6268 | -35.5707 | -36.5972
26.6329 | -35.5711 | -36.6023
26.639 | -35.5713 | -36.6077
26.6451 | -35.5717 | -36.6134
26.6512 | -35.572 | -36.6188
26.6573 | -35.5724 | -36.6243
26.6634 | -35.5728 | -36.6295
26.6695 | -35.5731 | -36.6352
26.6756 | -35.5734 | -36.6407
26.6817 | -35.5737 | -36.6463
26.6878 | -35.574 | -36.652
26.6939 | -35.5743 | -36.6573
26.7 | -35.5746 | -36.6621
26.7061 | -35.5749 | -36.6675
26.7122 | -35.5751 | -36.6727
26.7183 | -35.5754 | -36.6782
26.7245 | -35.5757 | -36.6832
26.7306 | -35.576 | -36.6884
26.7367 | -35.5764 | -36.6937
26.7428 | -35.5767 | -36.6989
26.7489 | -35.5769 | -36.7041
26.755 | -35.5772 | -36.7095
26.7611 | -35.5775 | -36.7152
26.7672 | -35.5778 | -36.7203
26.7733 | -35.5781 | -36.7256
26.7794 | -35.5784 | -36.7311
26.7855 | -35.5788 | -36.737
26.7916 | -35.579 | -36.7422
26.7977 | -35.5793 | -36.7474
26.8038 | -35.5797 | -36.7526
26.8099 | -35.58 | -36.7579
26.816 | -35.5803 | -36.7633
26.8221 | -35.5805 | -36.7688
26.8282 | -35.5808 | -36.7739
26.8343 | -35.5812 | -36.7792
26.8404 | -35.5816 | -36.785
26.8465 | -35.5818 | -36.7902
26.8526 | -35.582 | -36.7956
26.8587 | -35.5823 | -36.801
26.8649 | -35.5827 | -36.8064
26.871 | -35.583 | -36.8118
26.8771 | -35.5834 | -36.8168
26.8832 | -35.5837 | -36.8217
26.8893 | -35.584 | -36.827
26.8954 | -35.5843 | -36.8324
26.9015 | -35.5846 | -36.8379
26.9076 | -35.5849 | -36.8433
26.9137 | -35.5852 | -36.8481
26.9198 | -35.5855 | -36.8533
26.9259 | -35.5857 | -36.8588
26.932 | -35.586 | -36.8644
26.9381 | -35.5863 | -36.8695
26.9442 | -35.5867 | -36.8752
26.9503 | -35.587 | -36.8802
26.9564 | -35.5873 | -36.8852
26.9625 | -35.5876 | -36.8905
26.9686 | -35.5879 | -36.8953
26.9747 | -35.5882 | -36.9006
26.9808 | -35.5886 | -36.9058
26.9869 | -35.5889 | -36.9112
26.993 | -35.5892 | -36.9166
26.9991 | -35.5895 | -36.922
27.0052 | -35.5897 | -36.9276
27.0114 | -35.59 | -36.933
27.0175 | -35.5904 | -36.9381
27.0236 | -35.5907 | -36.9434
27.0297 | -35.591 | -36.9489
27.0358 | -35.5912 | -36.9544
27.0419 | -35.5915 | -36.9597
27.048 | -35.5918 | -36.9653
27.0541 | -35.5921 | -36.9702
27.0602 | -35.5924 | -36.9756
27.0663 | -35.5927 | -36.9806
27.0724 | -35.593 | -36.9855
27.0785 | -35.5934 | -36.9911
27.0846 | -35.5938 | -36.9961
27.0907 | -35.594 | -37.0023
27.0968 | -35.5944 | -37.0075
27.1029 | -35.5947 | -37.0127
27.109 | -35.595 | -37.0177
27.1151 | -35.5952 | -37.0232
27.1212 | -35.5955 | -37.0283
27.1273 | -35.5959 | -37.0339
27.1334 | -35.5962 | -37.0396
27.1395 | -35.5965 | -37.0452
27.1456 | -35.5968 | -37.0508
27.1518 | -35.597 | -37.056
27.1579 | -35.5973 | -37.0611
27.164 | -35.5976 | -37.0667
27.1701 | -35.598 | -37.0719
27.1762 | -35.5983 | -37.0771
27.1823 | -35.5986 | -37.0825
27.1884 | -35.5989 | -37.0882
27.1945 | -35.5992 | -37.0932
27.2006 | -35.5995 | -37.099
27.2067 | -35.5998 | -37.1044
27.2128 | -35.6002 | -37.1096
27.2189 | -35.6005 | -37.115
27.225 | -35.6009 | -37.1207
27.2311 | -35.6012 | -37.1259
27.2372 | -35.6015 | -37.1311
27.2433 | -35.6018 | -37.136
27.2494 | -35.6021 | -37.1415
27.2555 | -35.6024 | -37.146
27.2616 | -35.6026 | -37.1511
27.2677 | -35.6029 | -37.1565
27.2738 | -35.6032 | -37.1619
27.2799 | -35.6035 | -37.1675
27.286 | -35.6038 | -37.1725
27.2921 | -35.6041 | -37.1779
27.2983 | -35.6044 | -37.1828
27.3044 | -35.6046 | -37.1883
27.3105 | -35.6049 | -37.1941
27.3166 | -35.6052 | -37.1993
27.3227 | -35.6055 | -37.2048
27.3288 | -35.6059 | -37.21
27.3349 | -35.6061 | -37.2151
27.341 | -35.6064 | -37.2205
27.3471 | -35.6067 | -37.2258
27.3532 | -35.607 | -37.2309
27.3593 | -35.6072 | -37.2363
27.3654 | -35.6074 | -37.242
27.3715 | -35.6078 | -37.2472
27.3776 | -35.608 | -37.2525
27.3837 | -35.6083 | -37.258
27.3898 | -35.6085 | -37.2635
27.3959 | -35.6088 | -37.2687
27.402 | -35.6091 | -37.2728
27.4081 | -35.6094 | -37.2778
27.4142 | -35.6097 | -37.2833
27.4203 | -35.61 | -37.2884
27.4264 | -35.6104 | -37.2934
27.4325 | -35.6107 | -37.2984
27.4387 | -35.6109 | -37.304
27.4448 | -35.6113 | -37.309
27.4509 | -35.6115 | -37.314
27.457 | -35.6118 | -37.3191
27.4631 | -35.6121 | -37.3246
27.4692 | -35.6123 | -37.3297
27.4753 | -35.6126 | -37.3346
27.4814 | -35.6129 | -37.3399
27.4875 | -35.6132 | -37.3453
27.4936 | -35.6135 | -37.3503
27.4997 | -35.6139 | -37.3554
27.5058 | -35.6142 | -37.3601
27.5119 | -35.6145 | -37.3646
27.518 | -35.6148 | -37.3695
27.5241 | -35.6151 | -37.375
27.5302 | -35.6154 | -37.3805
27.5363 | -35.6156 | -37.3859
27.5424 | -35.616 | -37.3898
27.5485 | -35.6163 | -37.3952
27.5546 | -35.6167 | -37.4004
27.5607 | -35.617 | -37.4057
27.5668 | -35.6172 | -37.4108
27.5729 | -35.6175 | -37.4162
27.5791 | -35.6178 | -37.4214
27.5852 | -35.6181 | -37.4271
27.5913 | -35.6184 | -37.4321
27.5974 | -35.6188 | -37.4374
27.6035 | -35.6191 | -37.4429
27.6096 | -35.6195 | -37.448
27.6157 | -35.6198 | -37.4525
27.6218 | -35.6201 | -37.4582
27.6279 | -35.6203 | -37.4633
27.634 | -35.6206 | -37.4682
27.6401 | -35.621 | -37.4736
27.6462 | -35.6213 | -37.4787
27.6523 | -35.6217 | -37.4836
27.6584 | -35.6219 | -37.489
27.6645 | -35.6222 | -37.4947
27.6706 | -35.6225 | -37.5004
27.6767 | -35.6227 | -37.5057
27.6828 | -35.6228 | -37.5115
27.6889 | -35.6231 | -37.516
27.695 | -35.6234 | -37.5217
27.7011 | -35.6238 | -37.527
27.7072 | -35.6241 | -37.5328
27.7133 | -35.6244 | -37.538
27.7194 | -35.6246 | -37.5428
27.7256 | -35.6249 | -37.548
27.7317 | -35.6252 | -37.553
27.7378 | -35.6253 | -37.5584
27.7439 | -35.6257 | -37.5633
27.75 | -35.626 | -37.5682
27.7561 | -35.6261 | -37.5735
27.7622 | -35.6264 | -37.5788
27.7683 | -35.6266 | -37.584
27.7744 | -35.6269 | -37.5894
27.7805 | -35.6271 | -37.5943
27.7866 | -35.6273 | -37.5995
27.7927 | -35.6276 | -37.6052
27.7988 | -35.6279 | -37.6102
27.8049 | -35.6282 | -37.6152
27.811 | -35.6286 | -37.6199
27.8171 | -35.6288 | -37.6257
27.8232 | -35.629 | -37.6315
27.8293 | -35.6293 | -37.6362
27.8354 | -35.6295 | -37.6415
27.8415 | -35.6298 | -37.6469
27.8476 | -35.6301 | -37.6517
27.8537 | -35.6305 | -37.6564
27.8598 | -35.6307 | -37.6616
27.866 | -35.6309 | -37.6667
27.8721 | -35.6312 | -37.6719
27.8782 | -35.6314 | -37.6771
27.8843 | -35.6316 | -37.6826
27.8904 | -35.632 | -37.6876
27.8965 | -35.6323 | -37.6928
27.9026 | -35.6325 | -37.6983
27.9087 | -35.6328 | -37.7033
27.9148 | -35.6331 | -37.709
27.9209 | -35.6334 | -37.7139
27.927 | -35.6336 | -37.7186
27.9331 | -35.634 | -37.7239
27.9392 | -35.6343 | -37.7295
27.9453 | -35.6346 | -37.7345
27.9514 | -35.6349 | -37.739
27.9575 | -35.6352 | -37.7443
27.9636 | -35.6355 | -37.749
27.9697 | -35.6358 | -37.754
27.9758 | -35.6361 | -37.7596
27.9819 | -35.6364 | -37.7652
27.988 | -35.6366 | -37.77
27.9941 | -35.6368 | -37.7748
28.0002 | -35.637 | -37.7791
28.0063 | -35.6373 | -37.7843
28.0125 | -35.6376 | -37.789
28.0186 | -35.6379 | -37.7944
28.0247 | -35.6381 | -37.7998
28.0308 | -35.6384 | -37.8048
28.0369 | -35.6386 | -37.8101
28.043 | -35.6388 | -37.8155
28.0491 | -35.6391 | -37.8211
28.0552 | -35.6394 | -37.826
28.0613 | -35.6397 | -37.8308
28.0674 | -35.64 | -37.836
28.0735 | -35.6403 | -37.8413

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4669 | -42.8877
21.4908 | -35.5251 | -34.9723
16.4948 | -34.9191 | -27.5927
12.6602 | -34.5674 | -21.3244
9.7171 | -34.3466 | -16.4406
7.4582 | -34.2083 | -12.4634
5.7244 | -34.1022 | -9.2484
4.3936 | -34.044 | -6.652
3.3722 | -34.0099 | -4.6106
2.5883 | -33.9992 | -3.0572
1.9866 | -33.9943 | -1.8216
1.5248 | -33.9889 | -0.83089
1.1703 | -33.9985 | -0.18039
1 | -34.0052 | 0.1373
0.89822 | -33.9886 | 0.41011
0.68941 | -33.9707 | 0.99231
0.52919 | -33.9752 | 1.442
0.40612 | -33.9693 | 2.0465
0.31171 | -33.967 | 2.4956
0.23928 | -33.976 | 3.2173
0.18358 | -33.9723 | 3.6508
0.14091 | -33.9905 | 4.5339
0.10819 | -34.0088 | 5.6911
0.083017 | -34.0127 | 7.1752
0.063705 | -34.0692 | 9.0446
0.048923 | -34.1319 | 11.5428
0.037479 | -34.2414 | 14.7408
0.028801 | -34.4322 | 18.7467
0.022078 | -34.7491 | 24.0578
0.017023 | -35.234 | 30.4246
0.013018 | -36.038 | 38.0891
0.010014 | -37.1471 | 45.9839

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003688 | -6.1552
0.007377 | -26.1035
0.011065 | -34.0779
0.014753 | -39.2069
0.018442 | -41.7867
0.02213 | -44.7312
0.025818 | -46.7266
0.029507 | -48.5853
0.033195 | -49.6692
0.036883 | -51.5018
0.040572 | -52.8228
0.04426 | -54.2022
0.047948 | -55.1575
0.051637 | -55.7632
0.055325 | -56.8501
0.059013 | -57.3989
0.062702 | -58.7201
0.06639 | -59.1713
0.070078 | -59.2709
0.073767 | -60.1488
0.077455 | -61.0533
0.081143 | -61.7291
0.084832 | -61.4651
0.08852 | -62.4284
0.092208 | -63.645
0.095897 | -63.942
0.099585 | -63.7921
0.10327 | -64.0868
0.10696 | -65.9425
0.11065 | -65.7692
0.11434 | -65.355
0.11803 | -65.8247
0.12172 | -65.7131
0.1254 | -64.9434
0.12909 | -65.4704
0.13278 | -65.7454
0.13647 | -66.5804
0.14016 | -67.1624
0.14385 | -67.6404
0.14753 | -68.2187
0.15122 | -69.5489
0.15491 | -68.7926
0.1586 | -68.7649
0.16229 | -69.2927
0.16598 | -69.6948
0.16966 | -70.598
0.17335 | -70.9211
0.17704 | -70.395
0.18073 | -70.4925
0.18442 | -70.7968
0.1881 | -69.8634
0.19179 | -70.0069
0.19548 | -70.0191
0.19917 | -70.2391
0.20286 | -71.3225
0.20655 | -71.5647
0.21024 | -71.2896
0.21392 | -71.1497
0.21761 | -71.1115
0.2213 | -72.6414
0.22499 | -72.1863
0.22868 | -72.5265
0.23236 | -73.525
0.23605 | -72.9704
0.23974 | -74.1864
0.24343 | -74.0621
0.24712 | -73.8254
0.25081 | -73.699
0.2545 | -72.9822
0.25818 | -74.2348
0.26187 | -75.0829
0.26556 | -74.552
0.26925 | -75.2702
0.27294 | -75.1097
0.27663 | -74.8133
0.28031 | -74.576
0.284 | -74.3957
0.28769 | -74.5937
0.29138 | -74.9512
0.29507 | -73.9718
0.29875 | -74.5208
0.30244 | -75.1813
0.30613 | -75.0989
0.30982 | -75.566
0.31351 | -75.9868
0.3172 | -75.6926
0.32088 | -75.8195
0.32457 | -76.1428
0.32826 | -75.5279
0.33195 | -75.8602
0.33564 | -76.5351
0.33933 | -76.7658
0.34302 | -76.9545
0.3467 | -76.6694
0.35039 | -78.4053
0.35408 | -77.8985
0.35777 | -76.6468
0.36146 | -77.1878
0.36514 | -76.7776
0.36883 | -77.0887
0.37252 | -78.3542
0.37621 | -78.0708
0.3799 | -78.4087
0.38359 | -77.3711
0.38727 | -76.9932
0.39096 | -77.3872
0.39465 | -77.2862
0.39834 | -79.1969
0.40203 | -78.9116
0.40572 | -77.3327
0.40941 | -77.2985
0.41309 | -78.0135
0.41678 | -78.5644
0.42047 | -79.0324
0.42416 | -79.0986
0.42785 | -78.8309
0.43154 | -79.1139
0.43522 | -80.8311
0.43891 | -80.789
0.4426 | -79.8768
0.44629 | -79.3046
0.44998 | -79.2224
0.45366 | -79.3188
0.45735 | -79.2142
0.46104 | -78.8839
0.46473 | -79.4851
0.46842 | -79.3747
0.47211 | -78.6393
0.4758 | -79.0562
0.47948 | -79.8665
0.48317 | -79.9236
0.48686 | -79.6817
0.49055 | -79.8562
0.49424 | -80.6229
0.49793 | -80.3779
0.50161 | -80.1036
0.5053 | -80.2287
0.50899 | -80.1227
0.51268 | -80.0435
0.51637 | -80.7363
0.52006 | -81.0859
0.52374 | -81.3544
0.52743 | -80.4283
0.53112 | -80.1132
0.53481 | -80.3812
0.5385 | -80.9091
0.54219 | -81.4311
0.54587 | -81.6073
0.54956 | -81.3677
0.55325 | -81.4025
0.55694 | -82.0736
0.56063 | -82.7143
0.56432 | -82.6099
0.568 | -82.4245
0.57169 | -82.1446
0.57538 | -81.5545
0.57907 | -81.4136
0.58276 | -81.8468
0.58644 | -82.2703
0.59013 | -82.2728
0.59382 | -82.1957
0.59751 | -81.968
0.6012 | -82.0323
0.60489 | -82.1716
0.60857 | -81.9228
0.61226 | -81.7118
0.61595 | -81.6928
0.61964 | -81.8507
0.62333 | -81.8956
0.62702 | -81.6684
0.6307 | -81.1828
0.63439 | -81.1176
0.63808 | -81.6779
0.64177 | -82.5366
0.64546 | -82.9714
0.64915 | -83.0182
0.65284 | -82.9722
0.65652 | -82.7677
0.66021 | -82.7503
0.6639 | -82.8827
0.66759 | -82.8816
0.67128 | -82.7673
0.67497 | -82.6041
0.67865 | -82.3111
0.68234 | -82.4022
0.68603 | -83.3127
0.68972 | -84.3333
0.69341 | -84.1992
0.6971 | -83.6524
0.70078 | -83.3245
0.70447 | -83.4379
0.70816 | -83.6383
0.71185 | -83.7842
0.71554 | -83.7362
0.71923 | -83.5173
0.72291 | -83.3057
0.7266 | -83.3528
0.73029 | -83.7541
0.73398 | -84.0565
0.73767 | -83.825
0.74135 | -83.3576
0.74504 | -83.5704
0.74873 | -84.056
0.75242 | -84.0248
0.75611 | -83.8162
0.7598 | -83.8294
0.76348 | -83.8531
0.76717 | -83.8328
0.77086 | -83.9495
0.77455 | -84.4014
0.77824 | -84.9919
0.78193 | -84.971
0.78561 | -84.3834
0.7893 | -84.3358
0.79299 | -84.8849
0.79668 | -84.8823
0.80037 | -84.596
0.80406 | -84.5088
0.80775 | -84.2682
0.81143 | -83.915
0.81512 | -84.0279
0.81881 | -84.9045
0.8225 | -85.8675
0.82619 | -85.9373
0.82988 | -85.3486
0.83356 | -85.1291
0.83725 | -85.4542
0.84094 | -85.8391
0.84463 | -85.7521
0.84832 | -85.2774
0.85201 | -85.2455
0.85569 | -85.767
0.85938 | -85.5033
0.86307 | -84.7365
0.86676 | -84.5893
0.87045 | -85.1973
0.87414 | -85.8089
0.87782 | -85.6815
0.88151 | -85.4233
0.8852 | -85.5367
0.88889 | -85.848
0.89258 | -86.1195
0.89627 | -86.1681
0.89995 | -85.9287
0.90364 | -85.786
0.90733 | -85.5366
0.91102 | -85.2383
0.91471 | -85.2446
0.9184 | -85.212
0.92208 | -85.5406
0.92577 | -86.3982
0.92946 | -86.5807
0.93315 | -86.0684
0.93684 | -85.7858
0.94053 | -85.8356
0.94421 | -86.2652
0.9479 | -86.618
0.95159 | -86.4385
0.95528 | -86.3038
0.95897 | -86.437
0.96266 | -86.4218
0.96634 | -86.3982
0.97003 | -86.4904
0.97372 | -86.2538
0.97741 | -85.9241
0.9811 | -86.1045
0.98479 | -86.9754
0.98847 | -87.608
0.99216 | -87.2905
0.99585 | -86.7608
0.99954 | -86.3598
1.0032 | -86.1329
1.0069 | -86.1991
1.0106 | -86.324
1.0143 | -86.5669
1.018 | -86.8282
1.0217 | -86.6499
1.0254 | -86.1979
1.029 | -86.0727
1.0327 | -86.3031
1.0364 | -86.5245
1.0401 | -86.6569
1.0438 | -86.9628
1.0475 | -87.3055
1.0512 | -87.6112
1.0549 | -87.659
1.0586 | -87.6127
1.0622 | -87.4751
1.0659 | -87.3513
1.0696 | -87.2318
1.0733 | -87.132
1.077 | -87.148
1.0807 | -87.2934
1.0844 | -87.4072
1.0881 | -87.4539
1.0917 | -87.5396
1.0954 | -87.6204
1.0991 | -87.6282
1.1028 | -87.647
1.1065 | -87.735
1.1102 | -87.9042
1.1139 | -88.1888
1.1176 | -88.509
1.1213 | -88.6115
1.1249 | -88.4567
1.1286 | -88.2972
1.1323 | -88.3081
1.136 | -88.4399
1.1397 | -88.5098
1.1434 | -88.442
1.1471 | -88.2951
1.1508 | -88.0295
1.1544 | -87.6427
1.1581 | -87.3454
1.1618 | -87.3338
1.1655 | -87.5691
1.1692 | -87.8749
1.1729 | -88.1872
1.1766 | -88.4784
1.1803 | -88.6103
1.184 | -88.5736
1.1876 | -88.5066
1.1913 | -88.508
1.195 | -88.5048
1.1987 | -88.3932
1.2024 | -88.3592
1.2061 | -88.5558
1.2098 | -88.7935
1.2135 | -88.871
1.2172 | -88.7762
1.2208 | -88.6048
1.2245 | -88.475
1.2282 | -88.4682
1.2319 | -88.6228
1.2356 | -88.9399
1.2393 | -89.3423
1.243 | -89.6615
1.2467 | -89.6757
1.2503 | -89.3
1.254 | -88.7386
1.2577 | -88.3747
1.2614 | -88.3688
1.2651 | -88.5973
1.2688 | -88.8507
1.2725 | -89.0085
1.2762 | -89.1343
1.2799 | -89.3036
1.2835 | -89.3936
1.2872 | -89.275
1.2909 | -89.0345
1.2946 | -88.9151
1.2983 | -89.0214
1.302 | -89.2923
1.3057 | -89.5235
1.3094 | -89.4384
1.313 | -89.1368
1.3167 | -88.9721
1.3204 | -89.0231
1.3241 | -89.0403
1.3278 | -88.8566
1.3315 | -88.6005
1.3352 | -88.4192
1.3389 | -88.4129
1.3426 | -88.5941
1.3462 | -88.8924
1.3499 | -89.251
1.3536 | -89.5803
1.3573 | -89.7362
1.361 | -89.6793
1.3647 | -89.6911
1.3684 | -90.0602
1.3721 | -90.5602
1.3757 | -90.6283
1.3794 | -90.262
1.3831 | -90.0002
1.3868 | -90.0087
1.3905 | -90.0104
1.3942 | -89.7464
1.3979 | -89.2658
1.4016 | -88.9324
1.4053 | -89.0577
1.4089 | -89.6091
1.4126 | -90.1218
1.4163 | -90.142
1.42 | -89.8426
1.4237 | -89.5867
1.4274 | -89.5248
1.4311 | -89.6092
1.4348 | -89.5996
1.4385 | -89.4575
1.4421 | -89.3944
1.4458 | -89.5018
1.4495 | -89.6928
1.4532 | -89.8153
1.4569 | -89.8163
1.4606 | -89.8058
1.4643 | -89.8228
1.468 | -89.8704
1.4716 | -89.9611
1.4753 | -89.9917
1.479 | -89.9872
1.4827 | -90.084
1.4864 | -90.2521
1.4901 | -90.3082
1.4938 | -90.138
1.4975 | -89.9013
1.5012 | -89.8092
1.5048 | -89.9125
1.5085 | -90.1263
1.5122 | -90.3072
1.5159 | -90.3792
1.5196 | -90.371
1.5233 | -90.3609
1.527 | -90.3765
1.5307 | -90.494
1.5343 | -90.8417
1.538 | -91.17
1.5417 | -91.0366
1.5454 | -90.6358
1.5491 | -90.4749
1.5528 | -90.6412
1.5565 | -90.8266
1.5602 | -90.873
1.5639 | -90.9192
1.5675 | -90.9797
1.5712 | -90.9113
1.5749 | -90.6129
1.5786 | -90.2816
1.5823 | -90.3041
1.586 | -90.7468
1.5897 | -91.2614
1.5934 | -91.3364
1.597 | -91.402
1.6007 | -91.3584
1.6044 | -91.2361
1.6081 | -91.0669
1.6118 | -90.9268
1.6155 | -90.9175
1.6192 | -91.0672
1.6229 | -91.2873
1.6266 | -91.3817
1.6302 | -91.1866
1.6339 | -90.8366
1.6376 | -90.5951
1.6413 | -90.6199
1.645 | -90.8951
1.6487 | -91.2625
1.6524 | -91.5417
1.6561 | -91.6679
1.6598 | -91.6976
1.6634 | -91.6498
1.6671 | -91.4959
1.6708 | -91.3099
1.6745 | -91.1813
1.6782 | -91.1631
1.6819 | -91.2406
1.6856 | -91.3489
1.6893 | -91.4382
1.6929 | -91.4732
1.6966 | -91.4512
1.7003 | -91.3724
1.704 | -91.2552
1.7077 | -91.1762
1.7114 | -91.197
1.7151 | -91.3283
1.7188 | -91.5313
1.7225 | -91.7306
1.7261 | -91.828
1.7298 | -91.7811
1.7335 | -91.6642
1.7372 | -91.5474
1.7409 | -91.3828
1.7446 | -91.1276
1.7483 | -90.8664
1.752 | -90.71
1.7556 | -90.7031
1.7593 | -90.8455
1.763 | -91.1139
1.7667 | -91.4519
1.7704 | -91.7385
1.7741 | -91.8628
1.7778 | -91.8602
1.7815 | -91.8563
1.7852 | -91.9104
1.7888 | -91.9835
1.7925 | -92.0151
1.7962 | -91.9407
1.7999 | -91.7211
1.8036 | -91.4342
1.8073 | -91.2016
1.811 | -91.0873
1.8147 | -91.1342
1.8183 | -91.3872
1.822 | -91.8194
1.8257 | -92.2535
1.8294 | -92.5002
1.8331 | -92.5634
1.8368 | -92.5604
1.8405 | -92.5494
1.8442 | -92.501
1.8479 | -92.4218
1.8515 | -92.3603
1.8552 | -92.3018
1.8589 | -92.2275
1.8626 | -92.138
1.8663 | -92.0538
1.87 | -92.0073
1.8737 | -91.9973
1.8774 | -92.0125
1.8811 | -92.039
1.8847 | -92.1062
1.8884 | -92.2463
1.8921 | -92.4065
1.8958 | -92.4899
1.8995 | -92.4667
1.9032 | -92.3661
1.9069 | -92.2461
1.9106 | -92.1371
1.9142 | -92.0095
1.9179 | -91.8361
1.9216 | -91.6933
1.9253 | -91.6827
1.929 | -91.8577
1.9327 | -92.1736
1.9364 | -92.47
1.9401 | -92.5537
1.9438 | -92.41
1.9474 | -92.2367
1.9511 | -92.2051
1.9548 | -92.3671
1.9585 | -92.6602
1.9622 | -92.9188
1.9659 | -93.0049
1.9696 | -92.9338
1.9733 | -92.7952
1.9769 | -92.6278
1.9806 | -92.4244
1.9843 | -92.2066
1.988 | -92.0431
1.9917 | -92.0004
1.9954 | -92.1076
1.9991 | -92.3656
2.0028 | -92.7324
2.0065 | -93.0804
2.0101 | -93.2755
2.0138 | -93.2502
2.0175 | -93.083
2.0212 | -92.9674
2.0249 | -92.9822
2.0286 | -93.0424
2.0323 | -93.0522
2.036 | -93.0169
2.0396 | -92.9753
2.0433 | -92.9258
2.047 | -92.8425
2.0507 | -92.7204
2.0544 | -92.5576
2.0581 | -92.3875
2.0618 | -92.2569
2.0655 | -92.2396
2.0692 | -92.322
2.0728 | -92.4262
2.0765 | -92.4973
2.0802 | -92.5632
2.0839 | -92.5719
2.0876 | -92.4584
2.0913 | -92.2786
2.095 | -92.2041
2.0987 | -92.3409
2.1024 | -92.6627
2.106 | -92.9829
2.1097 | -93.1225
2.1134 | -93.0139
2.1171 | -92.7834
2.1208 | -92.6274
2.1245 | -92.6484
2.1282 | -92.8093
2.1319 | -92.9139
2.1355 | -92.9935
2.1392 | -93.0378
2.1429 | -93.0583
2.1466 | -93.0543
2.1503 | -93.0446
2.154 | -93.043
2.1577 | -93.054
2.1614 | -93.0816
2.1651 | -93.1133
2.1687 | -93.1571
2.1724 | -93.223
2.1761 | -93.2955
2.1798 | -93.3414
2.1835 | -93.3592
2.1872 | -93.3806
2.1909 | -93.4554
2.1946 | -93.5924
2.1982 | -93.7564
2.2019 | -93.8849
2.2056 | -93.9317
2.2093 | -93.8785
2.213 | -93.7744
2.2167 | -93.6758
2.2204 | -93.6389
2.2241 | -93.6803
2.2278 | -93.7832
2.2314 | -93.8793
2.2351 | -93.9089
2.2388 | -93.8454
2.2425 | -93.7225
2.2462 | -93.6033
2.2499 | -93.5301
2.2536 | -93.5059
2.2573 | -93.5267
2.2609 | -93.587
2.2646 | -93.6846
2.2683 | -93.8119
2.272 | -93.9592
2.2757 | -94.0935
2.2794 | -94.1803
2.2831 | -94.1867
2.2868 | -94.0826
2.2905 | -93.883
2.2941 | -93.6435
2.2978 | -93.4496
2.3015 | -93.3584
2.3052 | -93.387
2.3089 | -93.5068
2.3126 | -93.6753
2.3163 | -93.8352
2.32 | -93.9394
2.3237 | -93.9726
2.3273 | -93.9551
2.331 | -93.9141
2.3347 | -93.8645
2.3384 | -93.7905
2.3421 | -93.7016
2.3458 | -93.6402
2.3495 | -93.633
2.3532 | -93.6716
2.3568 | -93.7305
2.3605 | -93.7864
2.3642 | -93.81
2.3679 | -93.7988
2.3716 | -93.758
2.3753 | -93.6894
2.379 | -93.6055
2.3827 | -93.5374
2.3864 | -93.5219
2.39 | -93.5868
2.3937 | -93.7272
2.3974 | -93.8993
2.4011 | -94.0559
2.4048 | -94.1661
2.4085 | -94.2107
2.4122 | -94.1958
2.4159 | -94.1632
2.4195 | -94.1611
2.4232 | -94.2184
2.4269 | -94.3318
2.4306 | -94.4711
2.4343 | -94.5878
2.438 | -94.6539
2.4417 | -94.6728
2.4454 | -94.6885
2.4491 | -94.6978
2.4527 | -94.7064
2.4564 | -94.6897
2.4601 | -94.6287
2.4638 | -94.531
2.4675 | -94.4285
2.4712 | -94.3448
2.4749 | -94.2846
2.4786 | -94.2459
2.4823 | -94.2014
2.4859 | -94.1376
2.4896 | -94.0914
2.4933 | -94.0999
2.497 | -94.1913
2.5007 | -94.3384
2.5044 | -94.4611
2.5081 | -94.5004
2.5118 | -94.465
2.5154 | -94.41
2.5191 | -94.3905
2.5228 | -94.4411
2.5265 | -94.5555
2.5302 | -94.6891
2.5339 | -94.773
2.5376 | -94.742
2.5413 | -94.6041
2.545 | -94.4341
2.5486 | -94.3231
2.5523 | -94.3155
2.556 | -94.4167
2.5597 | -94.5946
2.5634 | -94.7934
2.5671 | -94.9374
2.5708 | -94.9662
2.5745 | -94.8808
2.5781 | -94.7289
2.5818 | -94.5596
2.5855 | -94.4171
2.5892 | -94.314
2.5929 | -94.242
2.5966 | -94.1958
2.6003 | -94.1746
2.604 | -94.1838
2.6077 | -94.235
2.6113 | -94.3512
2.615 | -94.5229
2.6187 | -94.7266
2.6224 | -94.89
2.6261 | -94.9563
2.6298 | -94.9213
2.6335 | -94.8285
2.6372 | -94.7137
2.6408 | -94.6185
2.6445 | -94.5685
2.6482 | -94.5731
2.6519 | -94.6304
2.6556 | -94.758
2.6593 | -94.9795
2.663 | -95.1645
2.6667 | -95.3246
2.6704 | -95.4005
2.674 | -95.3667
2.6777 | -95.238
2.6814 | -95.0582
2.6851 | -94.8771
2.6888 | -94.7419
2.6925 | -94.687
2.6962 | -94.7101
2.6999 | -94.7844
2.7035 | -94.8862
2.7072 | -94.9858
2.7109 | -95.0673
2.7146 | -95.1184
2.7183 | -95.1114
2.722 | -95.0474
2.7257 | -94.9494
2.7294 | -94.8678
2.7331 | -94.851
2.7367 | -94.9226
2.7404 | -95.0755
2.7441 | -95.2922
2.7478 | -95.5274
2.7515 | -95.7115
2.7552 | -95.7988
2.7589 | -95.7699
2.7626 | -95.6465
2.7663 | -95.4881
2.7699 | -95.3318
2.7736 | -95.2045
2.7773 | -95.1011
2.781 | -94.9976
2.7847 | -94.883
2.7884 | -94.7873
2.7921 | -94.7337
2.7958 | -94.7527
2.7994 | -94.8415
2.8031 | -94.9727
2.8068 | -95.1105
2.8105 | -95.2181
2.8142 | -95.2759
2.8179 | -95.3026
2.8216 | -95.3211
2.8253 | -95.3602
2.829 | -95.4379
2.8326 | -95.5545
2.8363 | -95.6904
2.84 | -95.828
2.8437 | -95.9351
2.8474 | -95.9921
2.8511 | -95.9707
2.8548 | -95.8741
2.8585 | -95.732
2.8621 | -95.5645
2.8658 | -95.3975
2.8695 | -95.2446
2.8732 | -95.1215
2.8769 | -95.0437
2.8806 | -95.0126
2.8843 | -95.0335
2.888 | -95.1058
2.8917 | -95.1945
2.8953 | -95.2546
2.899 | -95.2799
2.9027 | -95.2737
2.9064 | -95.2743
2.9101 | -95.3202
2.9138 | -95.4062
2.9175 | -95.5328
2.9212 | -95.6556
2.9249 | -95.7313
2.9285 | -95.7271
2.9322 | -95.6382
2.9359 | -95.5108
2.9396 | -95.4015
2.9433 | -95.3468
2.947 | -95.3466
2.9507 | -95.3991
2.9544 | -95.4715
2.958 | -95.5394
2.9617 | -95.5792
2.9654 | -95.5733
2.9691 | -95.5417
2.9728 | -95.5094
2.9765 | -95.5135
2.9802 | -95.555
2.9839 | -95.6157
2.9876 | -95.6751
2.9912 | -95.6899
2.9949 | -95.6356
2.9986 | -95.5191
3.0023 | -95.3812
3.006 | -95.2675
3.0097 | -95.2229
3.0134 | -95.2615
3.0171 | -95.3663
3.0207 | -95.5158
3.0244 | -95.6794
3.0281 | -95.8384
3.0318 | -95.9664
3.0355 | -96.0343
3.0392 | -96.046
3.0429 | -95.9995
3.0466 | -95.8992
3.0503 | -95.8002
3.0539 | -95.7401
3.0576 | -95.7293
3.0613 | -95.7807
3.065 | -95.8892
3.0687 | -96.0047
3.0724 | -96.0734
3.0761 | -96.0797
3.0798 | -96.03
3.0834 | -95.9506
3.0871 | -95.8732
3.0908 | -95.8046
3.0945 | -95.758
3.0982 | -95.73
3.1019 | -95.7354
3.1056 | -95.7768
3.1093 | -95.8282
3.113 | -95.8527
3.1166 | -95.8418
3.1203 | -95.8089
3.124 | -95.7641
3.1277 | -95.7375
3.1314 | -95.7489
3.1351 | -95.7928
3.1388 | -95.856
3.1425 | -95.9072
3.1462 | -95.9404
3.1498 | -95.985
3.1535 | -96.0676
3.1572 | -96.184
3.1609 | -96.3157
3.1646 | -96.4163
3.1683 | -96.4595
3.172 | -96.4609
3.1757 | -96.4513
3.1793 | -96.4591
3.183 | -96.4989
3.1867 | -96.5722
3.1904 | -96.6505
3.1941 | -96.7132
3.1978 | -96.7666
3.2015 | -96.7984
3.2052 | -96.8084
3.2089 | -96.7963
3.2125 | -96.7456
3.2162 | -96.6533
3.2199 | -96.541
3.2236 | -96.426
3.2273 | -96.3177
3.231 | -96.2319
3.2347 | -96.1775
3.2384 | -96.1519
3.242 | -96.146
3.2457 | -96.1536
3.2494 | -96.1601
3.2531 | -96.1533
3.2568 | -96.1528
3.2605 | -96.1624
3.2642 | -96.1883
3.2679 | -96.2365
3.2716 | -96.2867
3.2752 | -96.3187
3.2789 | -96.3415
3.2826 | -96.3553
3.2863 | -96.3706
3.29 | -96.3948
3.2937 | -96.4229
3.2974 | -96.4485
3.3011 | -96.4682
3.3047 | -96.4875
3.3084 | -96.5213
3.3121 | -96.5564
3.3158 | -96.5862
3.3195 | -96.6111
3.3232 | -96.6104
3.3269 | -96.5944
3.3306 | -96.5691
3.3343 | -96.5401
3.3379 | -96.5251
3.3416 | -96.5187
3.3453 | -96.5226
3.349 | -96.5528
3.3527 | -96.6062
3.3564 | -96.6695
3.3601 | -96.7425
3.3638 | -96.8217
3.3674 | -96.9022
3.3711 | -96.9721
3.3748 | -97.0239
3.3785 | -97.0393
3.3822 | -96.9958
3.3859 | -96.8955
3.3896 | -96.7492
3.3933 | -96.5795
3.397 | -96.4203
3.4006 | -96.3169
3.4043 | -96.2867
3.408 | -96.3475
3.4117 | -96.4895
3.4154 | -96.6759
3.4191 | -96.878
3.4228 | -97.0674
3.4265 | -97.2196
3.4302 | -97.3109
3.4338 | -97.3543
3.4375 | -97.3464
3.4412 | -97.2978
3.4449 | -97.2382
3.4486 | -97.1784
3.4523 | -97.1309
3.456 | -97.0873
3.4597 | -97.0646
3.4633 | -97.0688
3.467 | -97.1019
3.4707 | -97.1462
3.4744 | -97.1873
3.4781 | -97.2267
3.4818 | -97.256
3.4855 | -97.2815
3.4892 | -97.3108
3.4929 | -97.354
3.4965 | -97.4137
3.5002 | -97.4882
3.5039 | -97.5738
3.5076 | -97.6599
3.5113 | -97.731
3.515 | -97.765
3.5187 | -97.7496
3.5224 | -97.7057
3.526 | -97.6315
3.5297 | -97.5264
3.5334 | -97.3996
3.5371 | -97.2703
3.5408 | -97.1581
3.5445 | -97.0948
3.5482 | -97.0881
3.5519 | -97.137
3.5556 | -97.2259
3.5592 | -97.329
3.5629 | -97.4267
3.5666 | -97.5021
3.5703 | -97.5256
3.574 | -97.4903
3.5777 | -97.4047
3.5814 | -97.2961
3.5851 | -97.187
3.5888 | -97.0946
3.5924 | -97.0342
3.5961 | -97.0094
3.5998 | -97.0217
3.6035 | -97.0813
3.6072 | -97.1783
3.6109 | -97.3012
3.6146 | -97.4301
3.6183 | -97.5414
3.6219 | -97.6309
3.6256 | -97.6931
3.6293 | -97.7227
3.633 | -97.7115
3.6367 | -97.6784
3.6404 | -97.6363
3.6441 | -97.5947
3.6478 | -97.557
3.6515 | -97.5164
3.6551 | -97.4657
3.6588 | -97.4059
3.6625 | -97.3591
3.6662 | -97.3227
3.6699 | -97.2936
3.6736 | -97.2725
3.6773 | -97.2513
3.681 | -97.234
3.6846 | -97.2137
3.6883 | -97.1834
3.692 | -97.1399
3.6957 | -97.0993
3.6994 | -97.0718
3.7031 | -97.0694
3.7068 | -97.103
3.7105 | -97.1658
3.7142 | -97.26
3.7178 | -97.3679
3.7215 | -97.459
3.7252 | -97.545
3.7289 | -97.6264
3.7326 | -97.6903
3.7363 | -97.744
3.74 | -97.7925
3.7437 | -97.8369
3.7473 | -97.8888
3.751 | -97.9374
3.7547 | -97.9806
3.7584 | -98.0233
3.7621 | -98.0473
3.7658 | -98.0568
3.7695 | -98.0625
3.7732 | -98.0587
3.7769 | -98.0547
3.7805 | -98.0721
3.7842 | -98.1038
3.7879 | -98.1409
3.7916 | -98.181
3.7953 | -98.2268
3.799 | -98.2621
3.8027 | -98.2711
3.8064 | -98.2628
3.8101 | -98.241
3.8137 | -98.1984
3.8174 | -98.1405
3.8211 | -98.089
3.8248 | -98.0538
3.8285 | -98.0356
3.8322 | -98.0388
3.8359 | -98.0691
3.8396 | -98.1312
3.8432 | -98.2046
3.8469 | -98.285
3.8506 | -98.3606
3.8543 | -98.4068
3.858 | -98.4063
3.8617 | -98.3615
3.8654 | -98.2841
3.8691 | -98.1793
3.8728 | -98.0714
3.8764 | -97.9855
3.8801 | -97.9253
3.8838 | -97.896
3.8875 | -97.896
3.8912 | -97.9144
3.8949 | -97.9511
3.8986 | -97.9999
3.9023 | -98.0592
3.9059 | -98.128
3.9096 | -98.1927
3.9133 | -98.2456
3.917 | -98.283
3.9207 | -98.3001
3.9244 | -98.303
3.9281 | -98.2922
3.9318 | -98.2918
3.9355 | -98.2968
3.9391 | -98.315
3.9428 | -98.3452
3.9465 | -98.3818
3.9502 | -98.4186
3.9539 | -98.4403
3.9576 | -98.4387
3.9613 | -98.4189
3.965 | -98.3797
3.9686 | -98.3304
3.9723 | -98.2881
3.976 | -98.2572
3.9797 | -98.2452
3.9834 | -98.2461
3.9871 | -98.2514
3.9908 | -98.2561
3.9945 | -98.2433
3.9982 | -98.2222
4.0018 | -98.1897
4.0055 | -98.1541
4.0092 | -98.1214
4.0129 | -98.1008
4.0166 | -98.0979
4.0203 | -98.1115
4.024 | -98.1432
4.0277 | -98.1773
4.0314 | -98.2121
4.035 | -98.2329
4.0387 | -98.2344
4.0424 | -98.2207
4.0461 | -98.1863
4.0498 | -98.1302
4.0535 | -98.0674
4.0572 | -98.0065
4.0609 | -97.9591
4.0645 | -97.9315
4.0682 | -97.9181
4.0719 | -97.9206
4.0756 | -97.927
4.0793 | -97.9268
4.083 | -97.9158
4.0867 | -97.9065
4.0904 | -97.9169
4.0941 | -97.9459
4.0977 | -97.9974
4.1014 | -98.0682
4.1051 | -98.149
4.1088 | -98.233
4.1125 | -98.3016
4.1162 | -98.3498
4.1199 | -98.3769
4.1236 | -98.3812
4.1272 | -98.3708
4.1309 | -98.3537
4.1346 | -98.3327
4.1383 | -98.3164
4.142 | -98.3144
4.1457 | -98.3357
4.1494 | -98.3746
4.1531 | -98.4244
4.1568 | -98.4766
4.1604 | -98.5215
4.1641 | -98.5465
4.1678 | -98.5318
4.1715 | -98.4795
4.1752 | -98.406
4.1789 | -98.316
4.1826 | -98.23
4.1863 | -98.1724
4.1899 | -98.1617
4.1936 | -98.2067
4.1973 | -98.3059
4.201 | -98.4497
4.2047 | -98.6208
4.2084 | -98.7924
4.2121 | -98.9504
4.2158 | -99.0901
4.2195 | -99.2023
4.2231 | -99.2922
4.2268 | -99.3648
4.2305 | -99.4309
4.2342 | -99.484
4.2379 | -99.5058
4.2416 | -99.4958
4.2453 | -99.4382
4.249 | -99.3244
4.2527 | -99.1688
4.2563 | -98.9881
4.26 | -98.8416
4.2637 | -98.7109
4.2674 | -98.6033
4.2711 | -98.5257
4.2748 | -98.4839
4.2785 | -98.4725
4.2822 | -98.4858
4.2858 | -98.5203
4.2895 | -98.563
4.2932 | -98.6057
4.2969 | -98.6355
4.3006 | -98.6465
4.3043 | -98.6427
4.308 | -98.627
4.3117 | -98.6162
4.3154 | -98.606
4.319 | -98.6046
4.3227 | -98.6204
4.3264 | -98.6586
4.3301 | -98.7154
4.3338 | -98.791
4.3375 | -98.8774
4.3412 | -98.9625
4.3449 | -99.0258
4.3485 | -99.0712
4.3522 | -99.0928
4.3559 | -99.078
4.3596 | -99.0361
4.3633 | -98.9767
4.367 | -98.9087
4.3707 | -98.8404
4.3744 | -98.7709
4.3781 | -98.7035
4.3817 | -98.6461
4.3854 | -98.6112
4.3891 | -98.5831
4.3928 | -98.5635
4.3965 | -98.563
4.4002 | -98.5665
4.4039 | -98.5724
4.4076 | -98.5874
4.4112 | -98.6081
4.4149 | -98.6403
4.4186 | -98.6792
4.4223 | -98.7306
4.426 | -98.7962
4.4297 | -98.8631
4.4334 | -98.9283
4.4371 | -98.9936
4.4408 | -99.0457
4.4444 | -99.0881
4.4481 | -99.1206
4.4518 | -99.1398
4.4555 | -99.1482
4.4592 | -99.1603
4.4629 | -99.1718
4.4666 | -99.1731
4.4703 | -99.1791
4.474 | -99.1839
4.4776 | -99.1724
4.4813 | -99.1539
4.485 | -99.127
4.4887 | -99.0924
4.4924 | -99.06
4.4961 | -99.0398
4.4998 | -99.0419
4.5035 | -99.0581
4.5071 | -99.0952
4.5108 | -99.1565
4.5145 | -99.2336
4.5182 | -99.3223
4.5219 | -99.4176
4.5256 | -99.5136
4.5293 | -99.592
4.533 | -99.6471
4.5367 | -99.6751
4.5403 | -99.6711
4.544 | -99.635
4.5477 | -99.5831
4.5514 | -99.5217
4.5551 | -99.4589
4.5588 | -99.3855
4.5625 | -99.3137
4.5662 | -99.2466
4.5698 | -99.1843
4.5735 | -99.1372
4.5772 | -99.0942
4.5809 | -99.0623
4.5846 | -99.0471
4.5883 | -99.0432
4.592 | -99.0591
4.5957 | -99.0965
4.5994 | -99.1583
4.603 | -99.2399
4.6067 | -99.3381
4.6104 | -99.4346
4.6141 | -99.5231
4.6178 | -99.6019
4.6215 | -99.666
4.6252 | -99.7183
4.6289 | -99.7711
4.6326 | -99.8114
4.6362 | -99.8432
4.6399 | -99.8657
4.6436 | -99.877
4.6473 | -99.881
4.651 | -99.867
4.6547 | -99.8419
4.6584 | -99.8259
4.6621 | -99.8022
4.6657 | -99.7611
4.6694 | -99.7104
4.6731 | -99.6492
4.6768 | -99.5813
4.6805 | -99.5206
4.6842 | -99.4735
4.6879 | -99.4408
4.6916 | -99.4192
4.6953 | -99.4099
4.6989 | -99.4153
4.7026 | -99.4395
4.7063 | -99.4807
4.71 | -99.5411
4.7137 | -99.624
4.7174 | -99.721
4.7211 | -99.83
4.7248 | -99.922
4.7284 | -99.9817
4.7321 | -100.0131
4.7358 | -100.021
4.7395 | -100.0036
4.7432 | -99.968
4.7469 | -99.9305
4.7506 | -99.8899
4.7543 | -99.8493
4.758 | -99.8098
4.7616 | -99.7725
4.7653 | -99.7329
4.769 | -99.6872
4.7727 | -99.6361
4.7764 | -99.5855
4.7801 | -99.5229
4.7838 | -99.4483
4.7875 | -99.3558
4.7911 | -99.2826
4.7948 | -99.2063
4.7985 | -99.1398
4.8022 | -99.0838
4.8059 | -99.0532
4.8096 | -99.0484
4.8133 | -99.0628
4.817 | -99.1021
4.8207 | -99.166
4.8243 | -99.2484
4.828 | -99.34
4.8317 | -99.4327
4.8354 | -99.5165
4.8391 | -99.5731
4.8428 | -99.6035
4.8465 | -99.6052
4.8502 | -99.5768
4.8539 | -99.5405
4.8575 | -99.4949
4.8612 | -99.4522
4.8649 | -99.4202
4.8686 | -99.4074
4.8723 | -99.4189
4.876 | -99.4504
4.8797 | -99.5024
4.8834 | -99.5665
4.887 | -99.6442
4.8907 | -99.7263
4.8944 | -99.8007
4.8981 | -99.8629
4.9018 | -99.9167
4.9055 | -99.9497
4.9092 | -99.9722
4.9129 | -99.9886
4.9166 | -99.9976
4.9202 | -99.9984
4.9239 | -99.9913
4.9276 | -99.9935
4.9313 | -99.9938
4.935 | -99.9948
4.9387 | -99.9933
4.9424 | -99.9918
4.9461 | -99.9973
4.9497 | -99.9986
4.9534 | -99.9888
4.9571 | -99.9708
4.9608 | -99.94
4.9645 | -99.9015
4.9682 | -99.8694
4.9719 | -99.8459
4.9756 | -99.8296
4.9793 | -99.83
4.9829 | -99.8534
4.9866 | -99.8889
4.9903 | -99.9353
4.994 | -99.9958
4.9977 | -100.0581
5.0014 | -100.1167
5.0051 | -100.1571
5.0088 | -100.1907
5.0124 | -100.2143
5.0161 | -100.2242
5.0198 | -100.2271
5.0235 | -100.2276
5.0272 | -100.235
5.0309 | -100.2412
5.0346 | -100.2469
5.0383 | -100.2598
5.042 | -100.2727
5.0456 | -100.2836
5.0493 | -100.2977
5.053 | -100.3224
5.0567 | -100.3457
5.0604 | -100.3677
5.0641 | -100.3852
5.0678 | -100.3983
5.0715 | -100.4008
5.0751 | -100.3905
5.0788 | -100.3719
5.0825 | -100.3433
5.0862 | -100.3116
5.0899 | -100.2777
5.0936 | -100.2336
5.0973 | -100.1821
5.101 | -100.1267
5.1047 | -100.072
5.1083 | -100.0252
5.112 | -99.9785
5.1157 | -99.9407
5.1194 | -99.9241
5.1231 | -99.9272
5.1268 | -99.9494
5.1305 | -99.9894
5.1342 | -100.0495
5.1379 | -100.1203
5.1415 | -100.1963
5.1452 | -100.261
5.1489 | -100.3096
5.1526 | -100.3383
5.1563 | -100.339
5.16 | -100.3149
5.1637 | -100.267
5.1674 | -100.2018
5.171 | -100.1316
5.1747 | -100.0559
5.1784 | -99.99
5.1821 | -99.9372
5.1858 | -99.8985
5.1895 | -99.8867
5.1932 | -99.8889
5.1969 | -99.9019
5.2006 | -99.9253
5.2042 | -99.951
5.2079 | -99.9688
5.2116 | -99.98
5.2153 | -99.9848
5.219 | -99.9823
5.2227 | -99.9855
5.2264 | -99.9934
5.2301 | -100.0108
5.2337 | -100.0435
5.2374 | -100.1003
5.2411 | -100.1772
5.2448 | -100.278
5.2485 | -100.3935
5.2522 | -100.5121
5.2559 | -100.6197
5.2596 | -100.7134
5.2633 | -100.7926
5.2669 | -100.8406
5.2706 | -100.866
5.2743 | -100.8675
5.278 | -100.8502
5.2817 | -100.8218
5.2854 | -100.7941
5.2891 | -100.7609
5.2928 | -100.7296
5.2965 | -100.7058
5.3001 | -100.6823
5.3038 | -100.664
5.3075 | -100.6433
5.3112 | -100.6149
5.3149 | -100.5701
5.3186 | -100.5118
5.3223 | -100.4703
5.326 | -100.4241
5.3296 | -100.3767
5.3333 | -100.338
5.337 | -100.3045
5.3407 | -100.2724
5.3444 | -100.2437
5.3481 | -100.2227
5.3518 | -100.2123
5.3555 | -100.2178
5.3592 | -100.2278
5.3628 | -100.249
5.3665 | -100.2806
5.3702 | -100.3239
5.3739 | -100.3721
5.3776 | -100.4185
5.3813 | -100.4682
5.385 | -100.512
5.3887 | -100.5501
5.3923 | -100.5824
5.396 | -100.5998
5.3997 | -100.5942
5.4034 | -100.5648
5.4071 | -100.5257
5.4108 | -100.4877
5.4145 | -100.4496
5.4182 | -100.416
5.4219 | -100.3979
5.4255 | -100.3979
5.4292 | -100.4142
5.4329 | -100.4464
5.4366 | -100.4899
5.4403 | -100.5366
5.444 | -100.5934
5.4477 | -100.6428
5.4514 | -100.6756
5.455 | -100.6925
5.4587 | -100.6963
5.4624 | -100.6837
5.4661 | -100.6468
5.4698 | -100.6042
5.4735 | -100.5526
5.4772 | -100.4943
5.4809 | -100.4395
5.4846 | -100.3936
5.4882 | -100.3529
5.4919 | -100.3115
5.4956 | -100.2769
5.4993 | -100.2587
5.503 | -100.2555
5.5067 | -100.2652
5.5104 | -100.2848
5.5141 | -100.317
5.5178 | -100.3575
5.5214 | -100.4066
5.5251 | -100.4591
5.5288 | -100.5173
5.5325 | -100.5795
5.5362 | -100.641
5.5399 | -100.7066
5.5436 | -100.7722
5.5473 | -100.8435
5.5509 | -100.9159
5.5546 | -100.9791
5.5583 | -101.0432
5.562 | -101.0919
5.5657 | -101.1247
5.5694 | -101.1379
5.5731 | -101.1306
5.5768 | -101.1101
5.5805 | -101.0761
5.5841 | -101.0277
5.5878 | -100.9709
5.5915 | -100.9211
5.5952 | -100.8883
5.5989 | -100.8492
5.6026 | -100.8207
5.6063 | -100.8032
5.61 | -100.7932
5.6136 | -100.7815
5.6173 | -100.7756
5.621 | -100.7748
5.6247 | -100.7698
5.6284 | -100.7639
5.6321 | -100.7595
5.6358 | -100.7661
5.6395 | -100.7824
5.6432 | -100.8026
5.6468 | -100.844
5.6505 | -100.8901
5.6542 | -100.9377
5.6579 | -100.988
5.6616 | -101.036
5.6653 | -101.0761
5.669 | -101.1159
5.6727 | -101.1547
5.6763 | -101.1846
5.68 | -101.2065
5.6837 | -101.2292
5.6874 | -101.2382
5.6911 | -101.2308
5.6948 | -101.2162
5.6985 | -101.2031
5.7022 | -101.1799
5.7059 | -101.1495
5.7095 | -101.1244
5.7132 | -101.1053
5.7169 | -101.0879
5.7206 | -101.0602
5.7243 | -101.0345
5.728 | -101.0201
5.7317 | -101.0038
5.7354 | -100.9953
5.739 | -101.0015
5.7427 | -101.0066
5.7464 | -101.0112
5.7501 | -101.0189
5.7538 | -101.0149
5.7575 | -101.0085
5.7612 | -101.0039
5.7649 | -100.987
5.7686 | -100.9712
5.7722 | -100.9553
5.7759 | -100.9425
5.7796 | -100.9354
5.7833 | -100.9225
5.787 | -100.9217
5.7907 | -100.9323
5.7944 | -100.9354
5.7981 | -100.9296
5.8018 | -100.9191
5.8054 | -100.9066
5.8091 | -100.8816
5.8128 | -100.8424
5.8165 | -100.7998
5.8202 | -100.756
5.8239 | -100.7118
5.8276 | -100.6864
5.8313 | -100.6814
5.8349 | -100.6932
5.8386 | -100.7254
5.8423 | -100.7705
5.846 | -100.8209
5.8497 | -100.8664
5.8534 | -100.9025
5.8571 | -100.9308
5.8608 | -100.9473
5.8645 | -100.9467
5.8681 | -100.9243
5.8718 | -100.8841
5.8755 | -100.8262
5.8792 | -100.7517
5.8829 | -100.6728
5.8866 | -100.5962
5.8903 | -100.5219
5.894 | -100.4543
5.8976 | -100.3983
5.9013 | -100.3605
5.905 | -100.3425
5.9087 | -100.3441
5.9124 | -100.3666
5.9161 | -100.4077
5.9198 | -100.4612
5.9235 | -100.5263
5.9272 | -100.5948
5.9308 | -100.6625
5.9345 | -100.7274
5.9382 | -100.7903
5.9419 | -100.8446
5.9456 | -100.8913
5.9493 | -100.9307
5.953 | -100.9627
5.9567 | -101.0008
5.9604 | -101.0417
5.964 | -101.087
5.9677 | -101.1293
5.9714 | -101.1723
5.9751 | -101.2169
5.9788 | -101.2615
5.9825 | -101.309
5.9862 | -101.3567
5.9899 | -101.3987
5.9935 | -101.4424
5.9972 | -101.4802
6.0009 | -101.5059
6.0046 | -101.516
6.0083 | -101.5142
6.012 | -101.5061
6.0157 | -101.4853
6.0194 | -101.4659
6.0231 | -101.4478
6.0267 | -101.4274
6.0304 | -101.4096
6.0341 | -101.3914
6.0378 | -101.3767
6.0415 | -101.3662
6.0452 | -101.3662
6.0489 | -101.3638
6.0526 | -101.3687
6.0562 | -101.3741
6.0599 | -101.3864
6.0636 | -101.4028
6.0673 | -101.414
6.071 | -101.4279
6.0747 | -101.4366
6.0784 | -101.4381
6.0821 | -101.4352
6.0858 | -101.4165
6.0894 | -101.3944
6.0931 | -101.3716
6.0968 | -101.3495
6.1005 | -101.3289
6.1042 | -101.3123
6.1079 | -101.3082
6.1116 | -101.3107
6.1153 | -101.3212
6.1189 | -101.3426
6.1226 | -101.3721
6.1263 | -101.4125
6.13 | -101.456
6.1337 | -101.5029
6.1374 | -101.5497
6.1411 | -101.6005
6.1448 | -101.6507
6.1485 | -101.7032
6.1521 | -101.7491
6.1558 | -101.7915
6.1595 | -101.8244
6.1632 | -101.8438
6.1669 | -101.8488
6.1706 | -101.8378
6.1743 | -101.8171
6.178 | -101.7941
6.1817 | -101.7684
6.1853 | -101.7397
6.189 | -101.7079
6.1927 | -101.6789
6.1964 | -101.6584
6.2001 | -101.6415
6.2038 | -101.6313
6.2075 | -101.6306
6.2112 | -101.6292
6.2148 | -101.6266
6.2185 | -101.6205
6.2222 | -101.6089
6.2259 | -101.5902
6.2296 | -101.5695
6.2333 | -101.5417
6.237 | -101.5083
6.2407 | -101.4809
6.2444 | -101.4558
6.248 | -101.4336
6.2517 | -101.4182
6.2554 | -101.4054
6.2591 | -101.3989
6.2628 | -101.3982
6.2665 | -101.399
6.2702 | -101.4006
6.2739 | -101.4098
6.2775 | -101.4356
6.2812 | -101.4719
6.2849 | -101.515
6.2886 | -101.5667
6.2923 | -101.6274
6.296 | -101.687
6.2997 | -101.738
6.3034 | -101.7904
6.3071 | -101.8294
6.3107 | -101.8577
6.3144 | -101.878
6.3181 | -101.8837
6.3218 | -101.8821
6.3255 | -101.8839
6.3292 | -101.8809
6.3329 | -101.8707
6.3366 | -101.8639
6.3402 | -101.8609
6.3439 | -101.8677
6.3476 | -101.88
6.3513 | -101.8934
6.355 | -101.9027
6.3587 | -101.9028
6.3624 | -101.8993
6.3661 | -101.8841
6.3698 | -101.8593
6.3734 | -101.8323
6.3771 | -101.7983
6.3808 | -101.7559
6.3845 | -101.722
6.3882 | -101.6863
6.3919 | -101.6604
6.3956 | -101.6464
6.3993 | -101.638
6.403 | -101.6408
6.4066 | -101.6492
6.4103 | -101.6601
6.414 | -101.6783
6.4177 | -101.6942
6.4214 | -101.7031
6.4251 | -101.7134
6.4288 | -101.7135
6.4325 | -101.7065
6.4361 | -101.6972
6.4398 | -101.6818
6.4435 | -101.6672
6.4472 | -101.6574
6.4509 | -101.6527
6.4546 | -101.6543
6.4583 | -101.6605
6.462 | -101.6756
6.4657 | -101.6848
6.4693 | -101.6901
6.473 | -101.7006
6.4767 | -101.7085
6.4804 | -101.7055
6.4841 | -101.7009
6.4878 | -101.6869
6.4915 | -101.6606
6.4952 | -101.6264
6.4988 | -101.5958
6.5025 | -101.5678
6.5062 | -101.5467
6.5099 | -101.5336
6.5136 | -101.5283
6.5173 | -101.542
6.521 | -101.5716
6.5247 | -101.6138
6.5284 | -101.6651
6.532 | -101.7236
6.5357 | -101.7837
6.5394 | -101.8424
6.5431 | -101.8914
6.5468 | -101.9315
6.5505 | -101.9651
6.5542 | -101.9812
6.5579 | -101.9816
6.5615 | -101.9712
6.5652 | -101.9573
6.5689 | -101.9325
6.5726 | -101.8946
6.5763 | -101.8624
6.58 | -101.8403
6.5837 | -101.8211
6.5874 | -101.8089
6.5911 | -101.8048
6.5947 | -101.8117
6.5984 | -101.8243
6.6021 | -101.847
6.6058 | -101.8786
6.6095 | -101.9146
6.6132 | -101.9471
6.6169 | -101.9742
6.6206 | -102.0033
6.6243 | -102.0269
6.6279 | -102.0505
6.6316 | -102.0687
6.6353 | -102.0769
6.639 | -102.0808
6.6427 | -102.0841
6.6464 | -102.0799
6.6501 | -102.0695
6.6538 | -102.0561
6.6574 | -102.0375
6.6611 | -102.0135
6.6648 | -101.9855
6.6685 | -101.9645
6.6722 | -101.9432
6.6759 | -101.9193
6.6796 | -101.8967
6.6833 | -101.8757
6.687 | -101.861
6.6906 | -101.8543
6.6943 | -101.8517
6.698 | -101.853
6.7017 | -101.8567
6.7054 | -101.8637
6.7091 | -101.8752
6.7128 | -101.8972
6.7165 | -101.9215
6.7201 | -101.9444
6.7238 | -101.9707
6.7275 | -101.9988
6.7312 | -102.0231
6.7349 | -102.0457
6.7386 | -102.0688
6.7423 | -102.0984
6.746 | -102.1323
6.7497 | -102.1624
6.7533 | -102.1958
6.757 | -102.2319
6.7607 | -102.2659
6.7644 | -102.3009
6.7681 | -102.3296
6.7718 | -102.3607
6.7755 | -102.3932
6.7792 | -102.4302
6.7828 | -102.4738
6.7865 | -102.5159
6.7902 | -102.5541
6.7939 | -102.5909
6.7976 | -102.6234
6.8013 | -102.6432
6.805 | -102.6537
6.8087 | -102.6516
6.8124 | -102.6353
6.816 | -102.6049
6.8197 | -102.5612
6.8234 | -102.5158
6.8271 | -102.4642
6.8308 | -102.4097
6.8345 | -102.3603
6.8382 | -102.3117
6.8419 | -102.2692
6.8456 | -102.2293
6.8492 | -102.1956
6.8529 | -102.1669
6.8566 | -102.1416
6.8603 | -102.1188
6.864 | -102.1008
6.8677 | -102.0855
6.8714 | -102.0735
6.8751 | -102.0683
6.8787 | -102.0615
6.8824 | -102.057
6.8861 | -102.0578
6.8898 | -102.0644
6.8935 | -102.074
6.8972 | -102.0933
6.9009 | -102.1223
6.9046 | -102.152
6.9083 | -102.1855
6.9119 | -102.2219
6.9156 | -102.2552
6.9193 | -102.2843
6.923 | -102.3126
6.9267 | -102.3328
6.9304 | -102.3493
6.9341 | -102.3596
6.9378 | -102.3577
6.9414 | -102.3468
6.9451 | -102.3256
6.9488 | -102.3074
6.9525 | -102.2901
6.9562 | -102.2722
6.9599 | -102.2607
6.9636 | -102.2522
6.9673 | -102.2467
6.971 | -102.2496
6.9746 | -102.2573
6.9783 | -102.2682
6.982 | -102.2922
6.9857 | -102.321
6.9894 | -102.3497
6.9931 | -102.3817
6.9968 | -102.416
7.0005 | -102.4458
7.0041 | -102.4754
7.0078 | -102.5104
7.0115 | -102.5509
7.0152 | -102.5821
7.0189 | -102.6087
7.0226 | -102.6318
7.0263 | -102.653
7.03 | -102.665
7.0337 | -102.6672
7.0373 | -102.6585
7.041 | -102.6322
7.0447 | -102.6014
7.0484 | -102.5645
7.0521 | -102.5157
7.0558 | -102.4652
7.0595 | -102.4157
7.0632 | -102.3615
7.0669 | -102.3013
7.0705 | -102.2465
7.0742 | -102.1945
7.0779 | -102.1429
7.0816 | -102.1004
7.0853 | -102.0702
7.089 | -102.0457
7.0927 | -102.0369
7.0964 | -102.0402
7.1 | -102.0545
7.1037 | -102.0794
7.1074 | -102.1146
7.1111 | -102.1573
7.1148 | -102.1999
7.1185 | -102.2419
7.1222 | -102.2874
7.1259 | -102.3303
7.1296 | -102.3603
7.1332 | -102.3814
7.1369 | -102.3958
7.1406 | -102.3983
7.1443 | -102.4011
7.148 | -102.404
7.1517 | -102.3962
7.1554 | -102.3898
7.1591 | -102.3899
7.1627 | -102.3844
7.1664 | -102.383
7.1701 | -102.3961
7.1738 | -102.4207
7.1775 | -102.4513
7.1812 | -102.4856
7.1849 | -102.5177
7.1886 | -102.5468
7.1923 | -102.5752
7.1959 | -102.5993
7.1996 | -102.6143
7.2033 | -102.6223
7.207 | -102.6223
7.2107 | -102.621
7.2144 | -102.6156
7.2181 | -102.6041
7.2218 | -102.5869
7.2255 | -102.5664
7.2291 | -102.552
7.2328 | -102.535
7.2365 | -102.5173
7.2402 | -102.5038
7.2439 | -102.4896
7.2476 | -102.481
7.2513 | -102.4684
7.255 | -102.4594
7.2586 | -102.4558
7.2623 | -102.452
7.266 | -102.4521
7.2697 | -102.4609
7.2734 | -102.4686
7.2771 | -102.4751
7.2808 | -102.4801
7.2845 | -102.49
7.2882 | -102.5041
7.2918 | -102.5263
7.2955 | -102.5514
7.2992 | -102.5768
7.3029 | -102.6081
7.3066 | -102.6404
7.3103 | -102.6715
7.314 | -102.7014
7.3177 | -102.7295
7.3213 | -102.7605
7.325 | -102.7872
7.3287 | -102.8135
7.3324 | -102.832
7.3361 | -102.8357
7.3398 | -102.8355
7.3435 | -102.8305
7.3472 | -102.8167
7.3509 | -102.797
7.3545 | -102.77
7.3582 | -102.7436
7.3619 | -102.7133
7.3656 | -102.6833
7.3693 | -102.6548
7.373 | -102.6287
7.3767 | -102.6103
7.3804 | -102.597
7.384 | -102.5893
7.3877 | -102.5937
7.3914 | -102.6055
7.3951 | -102.6212
7.3988 | -102.6437
7.4025 | -102.672
7.4062 | -102.7002
7.4099 | -102.734
7.4136 | -102.7689
7.4172 | -102.7997
7.4209 | -102.8287
7.4246 | -102.8531
7.4283 | -102.8721
7.432 | -102.8879
7.4357 | -102.8997
7.4394 | -102.909
7.4431 | -102.9168
7.4467 | -102.9173
7.4504 | -102.9296
7.4541 | -102.942
7.4578 | -102.9505
7.4615 | -102.9618
7.4652 | -102.9707
7.4689 | -102.9801
7.4726 | -102.9955
7.4763 | -103.0076
7.4799 | -103.0231
7.4836 | -103.0464
7.4873 | -103.0635
7.491 | -103.0835
7.4947 | -103.0992
7.4984 | -103.116
7.5021 | -103.1305
7.5058 | -103.1366
7.5095 | -103.1343
7.5131 | -103.1289
7.5168 | -103.1201
7.5205 | -103.0963
7.5242 | -103.0715
7.5279 | -103.042
7.5316 | -103.0018
7.5353 | -102.955
7.539 | -102.9071
7.5426 | -102.861
7.5463 | -102.8134
7.55 | -102.7699
7.5537 | -102.7234
7.5574 | -102.6779
7.5611 | -102.6356
7.5648 | -102.5971
7.5685 | -102.568
7.5722 | -102.5426
7.5758 | -102.5264
7.5795 | -102.525
7.5832 | -102.5241
7.5869 | -102.5275
7.5906 | -102.5368
7.5943 | -102.548
7.598 | -102.5532
7.6017 | -102.5591
7.6053 | -102.5664
7.609 | -102.5763
7.6127 | -102.5807
7.6164 | -102.5882
7.6201 | -102.6038
7.6238 | -102.6219
7.6275 | -102.6402
7.6312 | -102.6612
7.6349 | -102.6948
7.6385 | -102.7348
7.6422 | -102.7828
7.6459 | -102.8335
7.6496 | -102.8898
7.6533 | -102.9444
7.657 | -102.995
7.6607 | -103.0458
7.6644 | -103.0903
7.6681 | -103.1331
7.6717 | -103.1679
7.6754 | -103.19
7.6791 | -103.1987
7.6828 | -103.1995
7.6865 | -103.2028
7.6902 | -103.1963
7.6939 | -103.1774
7.6976 | -103.1593
7.7012 | -103.138
7.7049 | -103.1188
7.7086 | -103.0971
7.7123 | -103.0788
7.716 | -103.0638
7.7197 | -103.0477
7.7234 | -103.0294
7.7271 | -103.0096
7.7308 | -102.9894
7.7344 | -102.9674
7.7381 | -102.9489
7.7418 | -102.9279
7.7455 | -102.9061
7.7492 | -102.8814
7.7529 | -102.8551
7.7566 | -102.8274
7.7603 | -102.8033
7.7639 | -102.7763
7.7676 | -102.7525
7.7713 | -102.7326
7.775 | -102.7067
7.7787 | -102.6851
7.7824 | -102.6722
7.7861 | -102.6619
7.7898 | -102.6537
7.7935 | -102.6563
7.7971 | -102.6588
7.8008 | -102.6582
7.8045 | -102.6619
7.8082 | -102.6673
7.8119 | -102.6746
7.8156 | -102.6794
7.8193 | -102.6803
7.823 | -102.6829
7.8266 | -102.684
7.8303 | -102.6875
7.834 | -102.6946
7.8377 | -102.707
7.8414 | -102.7262
7.8451 | -102.7501
7.8488 | -102.7805
7.8525 | -102.8113
7.8562 | -102.8394
7.8598 | -102.8624
7.8635 | -102.8817
7.8672 | -102.8968
7.8709 | -102.9118
7.8746 | -102.9177
7.8783 | -102.9179
7.882 | -102.9122
7.8857 | -102.8989
7.8894 | -102.8874
7.893 | -102.8653
7.8967 | -102.8407
7.9004 | -102.8239
7.9041 | -102.8116
7.9078 | -102.7987
7.9115 | -102.796
7.9152 | -102.7984
7.9189 | -102.8052
7.9225 | -102.8258
7.9262 | -102.8499
7.9299 | -102.8804
7.9336 | -102.9137
7.9373 | -102.9441
7.941 | -102.9718
7.9447 | -102.9973
7.9484 | -103.0217
7.9521 | -103.0492
7.9557 | -103.0768
7.9594 | -103.1005
7.9631 | -103.1215
7.9668 | -103.1436
7.9705 | -103.172
7.9742 | -103.1999
7.9779 | -103.2228
7.9816 | -103.2442
7.9852 | -103.2601
7.9889 | -103.2744
7.9926 | -103.2852
7.9963 | -103.2986
8 | -103.3119
8.0037 | -103.3226
8.0074 | -103.3296
8.0111 | -103.3413
8.0148 | -103.3495
8.0184 | -103.349
8.0221 | -103.3482
8.0258 | -103.3479
8.0295 | -103.3453
8.0332 | -103.3456
8.0369 | -103.3451
8.0406 | -103.3372
8.0443 | -103.3355
8.0479 | -103.3313
8.0516 | -103.3247
8.0553 | -103.3224
8.059 | -103.3219
8.0627 | -103.3161
8.0664 | -103.3081
8.0701 | -103.2966
8.0738 | -103.2823
8.0775 | -103.2681
8.0811 | -103.2511
8.0848 | -103.2322
8.0885 | -103.2117
8.0922 | -103.195
8.0959 | -103.1779
8.0996 | -103.1615
8.1033 | -103.1473
8.107 | -103.1317
8.1106 | -103.1172
8.1143 | -103.1131
8.118 | -103.1094
8.1217 | -103.1072
8.1254 | -103.1099
8.1291 | -103.1189
8.1328 | -103.1301
8.1365 | -103.1438
8.1402 | -103.1591
8.1438 | -103.1698
8.1475 | -103.179
8.1512 | -103.185
8.1549 | -103.1874
8.1586 | -103.1852
8.1623 | -103.1818
8.166 | -103.1747
8.1697 | -103.1671
8.1734 | -103.1609
8.177 | -103.1482
8.1807 | -103.1357
8.1844 | -103.1259
8.1881 | -103.1201
8.1918 | -103.121
8.1955 | -103.127
8.1992 | -103.1405
8.2029 | -103.161
8.2065 | -103.186
8.2102 | -103.2133
8.2139 | -103.2351
8.2176 | -103.2571
8.2213 | -103.286
8.225 | -103.3149
8.2287 | -103.3463
8.2324 | -103.3691
8.2361 | -103.3775
8.2397 | -103.3838
8.2434 | -103.3856
8.2471 | -103.3787
8.2508 | -103.3666
8.2545 | -103.3478
8.2582 | -103.3258
8.2619 | -103.302
8.2656 | -103.2779
8.2692 | -103.251
8.2729 | -103.2248
8.2766 | -103.2015
8.2803 | -103.1838
8.284 | -103.1688
8.2877 | -103.154
8.2914 | -103.1492
8.2951 | -103.1521
8.2988 | -103.1632
8.3024 | -103.1786
8.3061 | -103.1973
8.3098 | -103.2195
8.3135 | -103.2471
8.3172 | -103.2768
8.3209 | -103.3046
8.3246 | -103.3358
8.3283 | -103.365
8.332 | -103.3951
8.3356 | -103.429
8.3393 | -103.4603
8.343 | -103.4769
8.3467 | -103.4948
8.3504 | -103.5092
8.3541 | -103.5207
8.3578 | -103.5312
8.3615 | -103.528
8.3651 | -103.5163
8.3688 | -103.5047
8.3725 | -103.4836
8.3762 | -103.4538
8.3799 | -103.4185
8.3836 | -103.3802
8.3873 | -103.3409
8.391 | -103.304
8.3947 | -103.2677
8.3983 | -103.2306
8.402 | -103.1972
8.4057 | -103.1788
8.4094 | -103.1718
8.4131 | -103.1701
8.4168 | -103.1781
8.4205 | -103.1944
8.4242 | -103.2189
8.4278 | -103.2532
8.4315 | -103.2914
8.4352 | -103.3331
8.4389 | -103.3872
8.4426 | -103.4456
8.4463 | -103.5033
8.45 | -103.5615
8.4537 | -103.6203
8.4574 | -103.679
8.461 | -103.7338
8.4647 | -103.7841
8.4684 | -103.8227
8.4721 | -103.8606
8.4758 | -103.8867
8.4795 | -103.8997
8.4832 | -103.9024
8.4869 | -103.8938
8.4905 | -103.8744
8.4942 | -103.8479
8.4979 | -103.8195
8.5016 | -103.7804
8.5053 | -103.7365
8.509 | -103.6884
8.5127 | -103.6372
8.5164 | -103.5889
8.5201 | -103.5463
8.5237 | -103.511
8.5274 | -103.4773
8.5311 | -103.4369
8.5348 | -103.399
8.5385 | -103.3683
8.5422 | -103.3439
8.5459 | -103.3239
8.5496 | -103.3094
8.5533 | -103.2989
8.5569 | -103.2971
8.5606 | -103.2972
8.5643 | -103.2932
8.568 | -103.2975
8.5717 | -103.3027
8.5754 | -103.3131
8.5791 | -103.3267
8.5828 | -103.3441
8.5864 | -103.3657
8.5901 | -103.3941
8.5938 | -103.428
8.5975 | -103.4591
8.6012 | -103.4969
8.6049 | -103.5399
8.6086 | -103.5866
8.6123 | -103.6361
8.616 | -103.6783
8.6196 | -103.7206
8.6233 | -103.7645
8.627 | -103.8084
8.6307 | -103.8468
8.6344 | -103.8779
8.6381 | -103.9021
8.6418 | -103.9158
8.6455 | -103.9188
8.6491 | -103.9126
8.6528 | -103.8988
8.6565 | -103.8783
8.6602 | -103.8494
8.6639 | -103.818
8.6676 | -103.7808
8.6713 | -103.736
8.675 | -103.6894
8.6787 | -103.6367
8.6823 | -103.5876
8.686 | -103.5392
8.6897 | -103.4982
8.6934 | -103.463
8.6971 | -103.4257
8.7008 | -103.3951
8.7045 | -103.3682
8.7082 | -103.3484
8.7118 | -103.3331
8.7155 | -103.3278
8.7192 | -103.3287
8.7229 | -103.3338
8.7266 | -103.347
8.7303 | -103.3657
8.734 | -103.3863
8.7377 | -103.4094
8.7414 | -103.4355
8.745 | -103.4611
8.7487 | -103.4872
8.7524 | -103.5095
8.7561 | -103.5286
8.7598 | -103.5453
8.7635 | -103.5589
8.7672 | -103.5685
8.7709 | -103.5786
8.7745 | -103.5852
8.7782 | -103.588
8.7819 | -103.5822
8.7856 | -103.573
8.7893 | -103.5667
8.793 | -103.5527
8.7967 | -103.5338
8.8004 | -103.5133
8.8041 | -103.4931
8.8077 | -103.4725
8.8114 | -103.4422
8.8151 | -103.4112
8.8188 | -103.3823
8.8225 | -103.358
8.8262 | -103.3377
8.8299 | -103.317
8.8336 | -103.2976
8.8373 | -103.2768
8.8409 | -103.2638
8.8446 | -103.2568
8.8483 | -103.2571
8.852 | -103.2625
8.8557 | -103.2715
8.8594 | -103.2898
8.8631 | -103.3137
8.8668 | -103.342
8.8704 | -103.368
8.8741 | -103.3941
8.8778 | -103.422
8.8815 | -103.4545
8.8852 | -103.4862
8.8889 | -103.5169
8.8926 | -103.5463
8.8963 | -103.571
8.9 | -103.5945
8.9036 | -103.6129
8.9073 | -103.6258
8.911 | -103.6392
8.9147 | -103.6468
8.9184 | -103.6466
8.9221 | -103.6414
8.9258 | -103.6305
8.9295 | -103.6179
8.9331 | -103.6047
8.9368 | -103.5953
8.9405 | -103.5836
8.9442 | -103.5704
8.9479 | -103.5583
8.9516 | -103.5483
8.9553 | -103.5433
8.959 | -103.5452
8.9627 | -103.5457
8.9663 | -103.5512
8.97 | -103.558
8.9737 | -103.5635
8.9774 | -103.5673
8.9811 | -103.5729
8.9848 | -103.583
8.9885 | -103.5909
8.9922 | -103.5956
8.9959 | -103.6071
8.9995 | -103.624
9.0032 | -103.6471
9.0069 | -103.6676
9.0106 | -103.6896
9.0143 | -103.7145
9.018 | -103.7371
9.0217 | -103.7602
9.0254 | -103.7833
9.029 | -103.8002
9.0327 | -103.817
9.0364 | -103.8312
9.0401 | -103.8397
9.0438 | -103.8394
9.0475 | -103.8445
9.0512 | -103.8465
9.0549 | -103.848
9.0586 | -103.8482
9.0622 | -103.8439
9.0659 | -103.8409
9.0696 | -103.8393
9.0733 | -103.8431
9.077 | -103.8465
9.0807 | -103.8463
9.0844 | -103.8468
9.0881 | -103.8451
9.0917 | -103.8428
9.0954 | -103.8501
9.0991 | -103.8528
9.1028 | -103.8536
9.1065 | -103.8579
9.1102 | -103.8684
9.1139 | -103.8787
9.1176 | -103.8856
9.1213 | -103.8968
9.1249 | -103.9144
9.1286 | -103.9268
9.1323 | -103.9407
9.136 | -103.9562
9.1397 | -103.9718
9.1434 | -103.986
9.1471 | -104.0023
9.1508 | -104.0233
9.1544 | -104.0372
9.1581 | -104.0442
9.1618 | -104.0526
9.1655 | -104.0609
9.1692 | -104.0641
9.1729 | -104.0641
9.1766 | -104.057
9.1803 | -104.0468
9.184 | -104.0354
9.1876 | -104.0223
9.1913 | -104.0037
9.195 | -103.9823
9.1987 | -103.962
9.2024 | -103.9354
9.2061 | -103.9109
9.2098 | -103.8897
9.2135 | -103.8726
9.2172 | -103.8615
9.2208 | -103.8533
9.2245 | -103.8516
9.2282 | -103.8561
9.2319 | -103.8563
9.2356 | -103.8666
9.2393 | -103.886
9.243 | -103.9102
9.2467 | -103.9394
9.2503 | -103.9738
9.254 | -104.0123
9.2577 | -104.0487
9.2614 | -104.0846
9.2651 | -104.1193
9.2688 | -104.1514
9.2725 | -104.1793
9.2762 | -104.2038
9.2799 | -104.2242
9.2835 | -104.237
9.2872 | -104.2417
9.2909 | -104.2442
9.2946 | -104.24
9.2983 | -104.2241
9.302 | -104.1997
9.3057 | -104.1752
9.3094 | -104.147
9.313 | -104.12
9.3167 | -104.0902
9.3204 | -104.0637
9.3241 | -104.0381
9.3278 | -104.0174
9.3315 | -104.0021
9.3352 | -103.9914
9.3389 | -103.9864
9.3426 | -103.984
9.3462 | -103.9862
9.3499 | -103.9906
9.3536 | -103.998
9.3573 | -104.007
9.361 | -104.016
9.3647 | -104.0243
9.3684 | -104.0332
9.3721 | -104.0395
9.3757 | -104.0429
9.3794 | -104.0462
9.3831 | -104.0467
9.3868 | -104.0478
9.3905 | -104.043
9.3942 | -104.0372
9.3979 | -104.0322
9.4016 | -104.0255
9.4053 | -104.0171
9.4089 | -104.0083
9.4126 | -103.9994
9.4163 | -103.9904
9.42 | -103.9823
9.4237 | -103.9774
9.4274 | -103.9773
9.4311 | -103.9872
9.4348 | -103.9997
9.4385 | -104.0128
9.4421 | -104.0316
9.4458 | -104.0539
9.4495 | -104.0767
9.4532 | -104.1009
9.4569 | -104.1218
9.4606 | -104.142
9.4643 | -104.1583
9.468 | -104.1697
9.4716 | -104.179
9.4753 | -104.1825
9.479 | -104.1805
9.4827 | -104.1774
9.4864 | -104.1767
9.4901 | -104.175
9.4938 | -104.1714
9.4975 | -104.1685
9.5012 | -104.1686
9.5048 | -104.172
9.5085 | -104.1728
9.5122 | -104.1712
9.5159 | -104.1726
9.5196 | -104.18
9.5233 | -104.1855
9.527 | -104.1955
9.5307 | -104.2035
9.5343 | -104.2133
9.538 | -104.2231
9.5417 | -104.2307
9.5454 | -104.2371
9.5491 | -104.2454
9.5528 | -104.2557
9.5565 | -104.269
9.5602 | -104.2768
9.5639 | -104.2822
9.5675 | -104.2847
9.5712 | -104.2884
9.5749 | -104.2802
9.5786 | -104.2739
9.5823 | -104.265
9.586 | -104.2532
9.5897 | -104.2388
9.5934 | -104.2225
9.597 | -104.2052
9.6007 | -104.1909
9.6044 | -104.1773
9.6081 | -104.165
9.6118 | -104.1507
9.6155 | -104.143
9.6192 | -104.1376
9.6229 | -104.1322
9.6266 | -104.1287
9.6302 | -104.1284
9.6339 | -104.1318
9.6376 | -104.1355
9.6413 | -104.1412
9.645 | -104.151
9.6487 | -104.1671
9.6524 | -104.1834
9.6561 | -104.1993
9.6598 | -104.215
9.6634 | -104.2322
9.6671 | -104.2489
9.6708 | -104.2656
9.6745 | -104.2751
9.6782 | -104.2878
9.6819 | -104.3054
9.6856 | -104.3181
9.6893 | -104.3266
9.6929 | -104.3303
9.6966 | -104.3333
9.7003 | -104.3378
9.704 | -104.3372
9.7077 | -104.3344
9.7114 | -104.3341
9.7151 | -104.3376
9.7188 | -104.3442
9.7225 | -104.3568
9.7261 | -104.3639
9.7298 | -104.3748
9.7335 | -104.3872
9.7372 | -104.4036
9.7409 | -104.4236
9.7446 | -104.4411
9.7483 | -104.4598
9.752 | -104.4777
9.7556 | -104.4961
9.7593 | -104.514
9.763 | -104.5279
9.7667 | -104.5376
9.7704 | -104.5451
9.7741 | -104.5483
9.7778 | -104.5473
9.7815 | -104.54
9.7852 | -104.5273
9.7888 | -104.5136
9.7925 | -104.4921
9.7962 | -104.4726
9.7999 | -104.4517
9.8036 | -104.4278
9.8073 | -104.4054
9.811 | -104.3793
9.8147 | -104.3535
9.8183 | -104.334
9.822 | -104.312
9.8257 | -104.2888
9.8294 | -104.2725
9.8331 | -104.2584
9.8368 | -104.2499
9.8405 | -104.2443
9.8442 | -104.2405
9.8479 | -104.2442
9.8515 | -104.2515
9.8552 | -104.2596
9.8589 | -104.2697
9.8626 | -104.2798
9.8663 | -104.2931
9.87 | -104.3112
9.8737 | -104.3315
9.8774 | -104.3491
9.8811 | -104.3631
9.8847 | -104.3766
9.8884 | -104.3944
9.8921 | -104.4099
9.8958 | -104.4251
9.8995 | -104.4387
9.9032 | -104.4497
9.9069 | -104.4634
9.9106 | -104.4839
9.9142 | -104.5017
9.9179 | -104.5155
9.9216 | -104.5304
9.9253 | -104.548
9.929 | -104.5657
9.9327 | -104.583
9.9364 | -104.596
9.9401 | -104.6045
9.9438 | -104.6132
9.9474 | -104.6226
9.9511 | -104.6324
9.9548 | -104.6434
9.9585 | -104.6498
9.9622 | -104.6474
9.9659 | -104.6469
9.9696 | -104.6449
9.9733 | -104.6384
9.9769 | -104.6316
9.9806 | -104.6206
9.9843 | -104.6074
9.988 | -104.5965
9.9917 | -104.5846
9.9954 | -104.5665
9.9991 | -104.5482
10.0028 | -104.5307
10.0065 | -104.5153
10.0101 | -104.4951
10.0138 | -104.473
10.0175 | -104.4505
10.0212 | -104.4319
10.0249 | -104.4136
10.0286 | -104.3948
10.0323 | -104.3775
10.036 | -104.3644
10.0396 | -104.3534
10.0433 | -104.3436
10.047 | -104.3334
10.0507 | -104.3258
10.0544 | -104.3166
10.0581 | -104.3111
10.0618 | -104.3136
10.0655 | -104.3187
10.0692 | -104.3248
10.0728 | -104.3281
10.0765 | -104.3306
10.0802 | -104.3448
10.0839 | -104.364
10.0876 | -104.3796
10.0913 | -104.396
10.095 | -104.414
10.0987 | -104.4349
10.1024 | -104.456
10.106 | -104.4813
10.1097 | -104.5036
10.1134 | -104.5248
10.1171 | -104.544
10.1208 | -104.5623
10.1245 | -104.5775
10.1282 | -104.5924
10.1319 | -104.6
10.1355 | -104.6013
10.1392 | -104.5962
10.1429 | -104.5932
10.1466 | -104.5871
10.1503 | -104.5742
10.154 | -104.5566
10.1577 | -104.5407
10.1614 | -104.5212
10.1651 | -104.4941
10.1687 | -104.466
10.1724 | -104.4346
10.1761 | -104.404
10.1798 | -104.3732
10.1835 | -104.3416
10.1872 | -104.3144
10.1909 | -104.2907
10.1946 | -104.2683
10.1982 | -104.2476
10.2019 | -104.2304
10.2056 | -104.2167
10.2093 | -104.2068
10.213 | -104.2033
10.2167 | -104.2049
10.2204 | -104.2163
10.2241 | -104.2333
10.2278 | -104.2457
10.2314 | -104.2566
10.2351 | -104.2755
10.2388 | -104.2977
10.2425 | -104.32
10.2462 | -104.3441
10.2499 | -104.3696
10.2536 | -104.3933
10.2573 | -104.4144
10.2609 | -104.432
10.2646 | -104.4468
10.2683 | -104.4628
10.272 | -104.4765
10.2757 | -104.4877
10.2794 | -104.4953
10.2831 | -104.4974
10.2868 | -104.4938
10.2905 | -104.4898
10.2941 | -104.4756
10.2978 | -104.4568
10.3015 | -104.4387
10.3052 | -104.416
10.3089 | -104.3876
10.3126 | -104.3585
10.3163 | -104.3339
10.32 | -104.3071
10.3237 | -104.2794
10.3273 | -104.25
10.331 | -104.2239
10.3347 | -104.203
10.3384 | -104.1858
10.3421 | -104.1739
10.3458 | -104.1696
10.3495 | -104.1705
10.3532 | -104.1782
10.3568 | -104.1928
10.3605 | -104.2134
10.3642 | -104.2368
10.3679 | -104.2641
10.3716 | -104.2949
10.3753 | -104.3309
10.379 | -104.3695
10.3827 | -104.4079
10.3864 | -104.4438
10.39 | -104.4763
10.3937 | -104.5073
10.3974 | -104.5353
10.4011 | -104.563
10.4048 | -104.5842
10.4085 | -104.6018
10.4122 | -104.6143
10.4159 | -104.623
10.4195 | -104.6246
10.4232 | -104.618
10.4269 | -104.6131
10.4306 | -104.6039
10.4343 | -104.5937
10.438 | -104.584
10.4417 | -104.5699
10.4454 | -104.5599
10.4491 | -104.5487
10.4527 | -104.5385
10.4564 | -104.5315
10.4601 | -104.5307
10.4638 | -104.5299
10.4675 | -104.5307
10.4712 | -104.5359
10.4749 | -104.5449
10.4786 | -104.5555
10.4823 | -104.5683
10.4859 | -104.5814
10.4896 | -104.5959
10.4933 | -104.6133
10.497 | -104.6331
10.5007 | -104.6517
10.5044 | -104.6694
10.5081 | -104.6837
10.5118 | -104.6989
10.5154 | -104.7147
10.5191 | -104.7287
10.5228 | -104.7375
10.5265 | -104.7418
10.5302 | -104.7503
10.5339 | -104.7576
10.5376 | -104.7529
10.5413 | -104.7467
10.545 | -104.7409
10.5486 | -104.7316
10.5523 | -104.7182
10.556 | -104.7098
10.5597 | -104.7041
10.5634 | -104.6998
10.5671 | -104.6977
10.5708 | -104.6998
10.5745 | -104.7066
10.5781 | -104.7079
10.5818 | -104.7117
10.5855 | -104.7224
10.5892 | -104.7328
10.5929 | -104.7427
10.5966 | -104.7537
10.6003 | -104.7646
10.604 | -104.7746
10.6077 | -104.7833
10.6113 | -104.7916
10.615 | -104.8005
10.6187 | -104.8032
10.6224 | -104.8042
10.6261 | -104.8048
10.6298 | -104.8029
10.6335 | -104.7923
10.6372 | -104.7858
10.6408 | -104.7637
10.6445 | -104.7484
10.6482 | -104.729
10.6519 | -104.7072
10.6556 | -104.6839
10.6593 | -104.6587
10.663 | -104.6343
10.6667 | -104.6042
10.6704 | -104.5741
10.674 | -104.5459
10.6777 | -104.522
10.6814 | -104.503
10.6851 | -104.4825
10.6888 | -104.4619
10.6925 | -104.4452
10.6962 | -104.4332
10.6999 | -104.4213
10.7035 | -104.412
10.7072 | -104.4075
10.7109 | -104.4023
10.7146 | -104.3995
10.7183 | -104.4015
10.722 | -104.4017
10.7257 | -104.4035
10.7294 | -104.4055
10.7331 | -104.4097
10.7367 | -104.4146
10.7404 | -104.413
10.7441 | -104.4102
10.7478 | -104.4063
10.7515 | -104.4017
10.7552 | -104.3967
10.7589 | -104.3902
10.7626 | -104.3864
10.7663 | -104.3775
10.7699 | -104.3685
10.7736 | -104.3627
10.7773 | -104.3597
10.781 | -104.3578
10.7847 | -104.3547
10.7884 | -104.353
10.7921 | -104.3491
10.7958 | -104.349
10.7994 | -104.3522
10.8031 | -104.3591
10.8068 | -104.37
10.8105 | -104.3828
10.8142 | -104.3997
10.8179 | -104.4198
10.8216 | -104.4413
10.8253 | -104.4662
10.829 | -104.4895
10.8326 | -104.5142
10.8363 | -104.5394
10.84 | -104.5645
10.8437 | -104.5936
10.8474 | -104.625
10.8511 | -104.6617
10.8548 | -104.6955
10.8585 | -104.7283
10.8621 | -104.7633
10.8658 | -104.7964
10.8695 | -104.8306
10.8732 | -104.8588
10.8769 | -104.8853
10.8806 | -104.9119
10.8843 | -104.9374
10.888 | -104.9622
10.8917 | -104.9822
10.8953 | -104.9974
10.899 | -105.0082
10.9027 | -105.0168
10.9064 | -105.0259
10.9101 | -105.0281
10.9138 | -105.0241
10.9175 | -105.0163
10.9212 | -105.0051
10.9248 | -104.997
10.9285 | -104.9859
10.9322 | -104.9691
10.9359 | -104.9489
10.9396 | -104.9292
10.9433 | -104.913
10.947 | -104.8971
10.9507 | -104.8762
10.9544 | -104.8563
10.958 | -104.8429
10.9617 | -104.8278
10.9654 | -104.8098
10.9691 | -104.7985
10.9728 | -104.7915
10.9765 | -104.7879
10.9802 | -104.7892
10.9839 | -104.7959
10.9876 | -104.809
10.9912 | -104.827
10.9949 | -104.8448
10.9986 | -104.8661
11.0023 | -104.8905
11.006 | -104.9172
11.0097 | -104.9387
11.0134 | -104.9577
11.0171 | -104.9774
11.0207 | -104.9957
11.0244 | -105.0144
11.0281 | -105.0302
11.0318 | -105.0473
11.0355 | -105.0589
11.0392 | -105.0657
11.0429 | -105.0702
11.0466 | -105.0719
11.0503 | -105.0725
11.0539 | -105.0709
11.0576 | -105.071
11.0613 | -105.0664
11.065 | -105.0575
11.0687 | -105.049
11.0724 | -105.0413
11.0761 | -105.0326
11.0798 | -105.0218
11.0834 | -105.0062
11.0871 | -104.9934
11.0908 | -104.9819
11.0945 | -104.9667
11.0982 | -104.9485
11.1019 | -104.9299
11.1056 | -104.9122
11.1093 | -104.9003
11.113 | -104.8876
11.1166 | -104.8714
11.1203 | -104.8591
11.124 | -104.8469
11.1277 | -104.8322
11.1314 | -104.817
11.1351 | -104.8011
11.1388 | -104.787
11.1425 | -104.7699
11.1462 | -104.7493
11.1498 | -104.7293
11.1535 | -104.7104
11.1572 | -104.6932
11.1609 | -104.6777
11.1646 | -104.6642
11.1683 | -104.6563
11.172 | -104.6555
11.1757 | -104.6494
11.1793 | -104.6424
11.183 | -104.6358
11.1867 | -104.6341
11.1904 | -104.6325
11.1941 | -104.6291
11.1978 | -104.6273
11.2015 | -104.6279
11.2052 | -104.6283
11.2089 | -104.6305
11.2125 | -104.6291
11.2162 | -104.6282
11.2199 | -104.6245
11.2236 | -104.6181
11.2273 | -104.6113
11.231 | -104.6032
11.2347 | -104.5925
11.2384 | -104.5796
11.242 | -104.5677
11.2457 | -104.5568
11.2494 | -104.5458
11.2531 | -104.5379
11.2568 | -104.5323
11.2605 | -104.5264
11.2642 | -104.5222
11.2679 | -104.52
11.2716 | -104.5127
11.2752 | -104.5067
11.2789 | -104.5067
11.2826 | -104.5077
11.2863 | -104.5088
11.29 | -104.5078
11.2937 | -104.5088
11.2974 | -104.5129
11.3011 | -104.5164
11.3047 | -104.5247
11.3084 | -104.5353
11.3121 | -104.5429
11.3158 | -104.5477
11.3195 | -104.5465
11.3232 | -104.5459
11.3269 | -104.5472
11.3306 | -104.5445
11.3343 | -104.5438
11.3379 | -104.5462
11.3416 | -104.5455
11.3453 | -104.5437
11.349 | -104.5428
11.3527 | -104.5401
11.3564 | -104.5397
11.3601 | -104.54
11.3638 | -104.5401
11.3674 | -104.5413
11.3711 | -104.5425
11.3748 | -104.5462
11.3785 | -104.554
11.3822 | -104.5627
11.3859 | -104.5751
11.3896 | -104.5908
11.3933 | -104.6074
11.397 | -104.6248
11.4006 | -104.6446
11.4043 | -104.6649
11.408 | -104.6838
11.4117 | -104.7054
11.4154 | -104.7284
11.4191 | -104.7484
11.4228 | -104.7691
11.4265 | -104.79
11.4302 | -104.8069
11.4338 | -104.818
11.4375 | -104.827
11.4412 | -104.8327
11.4449 | -104.831
11.4486 | -104.8229
11.4523 | -104.8147
11.456 | -104.8046
11.4597 | -104.7883
11.4633 | -104.77
11.467 | -104.7481
11.4707 | -104.72
11.4744 | -104.6947
11.4781 | -104.6733
11.4818 | -104.6559
11.4855 | -104.6423
11.4892 | -104.6273
11.4929 | -104.614
11.4965 | -104.6047
11.5002 | -104.5965
11.5039 | -104.5879
11.5076 | -104.5858
11.5113 | -104.5859
11.515 | -104.5888
11.5187 | -104.591
11.5224 | -104.5952
11.526 | -104.6053
11.5297 | -104.6186
11.5334 | -104.6316
11.5371 | -104.6466
11.5408 | -104.6621
11.5445 | -104.6811
11.5482 | -104.701
11.5519 | -104.7208
11.5556 | -104.741
11.5592 | -104.7646
11.5629 | -104.7911
11.5666 | -104.8176
11.5703 | -104.8426
11.574 | -104.8638
11.5777 | -104.8817
11.5814 | -104.8961
11.5851 | -104.9108
11.5888 | -104.9204
11.5924 | -104.9277
11.5961 | -104.9362
11.5998 | -104.9425
11.6035 | -104.9446
11.6072 | -104.9439
11.6109 | -104.9408
11.6146 | -104.9366
11.6183 | -104.9288
11.6219 | -104.9169
11.6256 | -104.9066
11.6293 | -104.8977
11.633 | -104.8922
11.6367 | -104.8902
11.6404 | -104.8849
11.6441 | -104.8779
11.6478 | -104.8766
11.6515 | -104.8748
11.6551 | -104.8733
11.6588 | -104.8754
11.6625 | -104.8823
11.6662 | -104.8903
11.6699 | -104.9037
11.6736 | -104.9213
11.6773 | -104.9424
11.681 | -104.9597
11.6846 | -104.9799
11.6883 | -105.0001
11.692 | -105.0202
11.6957 | -105.0415
11.6994 | -105.0642
11.7031 | -105.0774
11.7068 | -105.0954
11.7105 | -105.1119
11.7142 | -105.1268
11.7178 | -105.1404
11.7215 | -105.1507
11.7252 | -105.1602
11.7289 | -105.1646
11.7326 | -105.1673
11.7363 | -105.1705
11.74 | -105.1738
11.7437 | -105.1747
11.7473 | -105.1766
11.751 | -105.1774
11.7547 | -105.1795
11.7584 | -105.183
11.7621 | -105.1843
11.7658 | -105.1811
11.7695 | -105.1799
11.7732 | -105.1835
11.7769 | -105.1839
11.7805 | -105.183
11.7842 | -105.1828
11.7879 | -105.1809
11.7916 | -105.1776
11.7953 | -105.169
11.799 | -105.1543
11.8027 | -105.1384
11.8064 | -105.1235
11.8101 | -105.1071
11.8137 | -105.0879
11.8174 | -105.0657
11.8211 | -105.0478
11.8248 | -105.0296
11.8285 | -105.0042
11.8322 | -104.976
11.8359 | -104.9502
11.8396 | -104.925
11.8432 | -104.9024
11.8469 | -104.8812
11.8506 | -104.8648
11.8543 | -104.8513
11.858 | -104.8439
11.8617 | -104.8359
11.8654 | -104.826
11.8691 | -104.8235
11.8728 | -104.8293
11.8764 | -104.8405
11.8801 | -104.8535
11.8838 | -104.8718
11.8875 | -104.8904
11.8912 | -104.9108
11.8949 | -104.9342
11.8986 | -104.9594
11.9023 | -104.9866
11.9059 | -105.0126
11.9096 | -105.0376
11.9133 | -105.0595
11.917 | -105.0772
11.9207 | -105.0902
11.9244 | -105.1038
11.9281 | -105.1121
11.9318 | -105.1198
11.9355 | -105.1258
11.9391 | -105.128
11.9428 | -105.126
11.9465 | -105.1185
11.9502 | -105.1066
11.9539 | -105.0933
11.9576 | -105.0785
11.9613 | -105.0625
11.965 | -105.0441
11.9686 | -105.0277
11.9723 | -105.0087
11.976 | -104.989
11.9797 | -104.969
11.9834 | -104.9489
11.9871 | -104.9316
11.9908 | -104.9135
11.9945 | -104.8932
11.9982 | -104.8775
12.0018 | -104.8676
12.0055 | -104.858
12.0092 | -104.8484
12.0129 | -104.8393
12.0166 | -104.8304
12.0203 | -104.8233
12.024 | -104.8176
12.0277 | -104.814
12.0314 | -104.8096
12.035 | -104.808
12.0387 | -104.8093
12.0424 | -104.8052
12.0461 | -104.7996
12.0498 | -104.7953
12.0535 | -104.7932
12.0572 | -104.79
12.0609 | -104.7871
12.0645 | -104.7873
12.0682 | -104.7895
12.0719 | -104.7926
12.0756 | -104.7955
12.0793 | -104.7974
12.083 | -104.8
12.0867 | -104.808
12.0904 | -104.8207
12.0941 | -104.8296
12.0977 | -104.8365
12.1014 | -104.8447
12.1051 | -104.8551
12.1088 | -104.8683
12.1125 | -104.8843
12.1162 | -104.8989
12.1199 | -104.9086
12.1236 | -104.9227
12.1272 | -104.939
12.1309 | -104.9531
12.1346 | -104.968
12.1383 | -104.9836
12.142 | -104.9991
12.1457 | -105.0163
12.1494 | -105.0305
12.1531 | -105.0439
12.1568 | -105.0586
12.1604 | -105.0696
12.1641 | -105.0779
12.1678 | -105.0882
12.1715 | -105.0942
12.1752 | -105.0981
12.1789 | -105.1017
12.1826 | -105.1028
12.1863 | -105.0994
12.1899 | -105.0929
12.1936 | -105.0837
12.1973 | -105.0752
12.201 | -105.0633
12.2047 | -105.0485
12.2084 | -105.0347
12.2121 | -105.0169
12.2158 | -104.9936
12.2195 | -104.9715
12.2231 | -104.9489
12.2268 | -104.9261
12.2305 | -104.9065
12.2342 | -104.8868
12.2379 | -104.8702
12.2416 | -104.8515
12.2453 | -104.8333
12.249 | -104.815
12.2527 | -104.8012
12.2563 | -104.7917
12.26 | -104.782
12.2637 | -104.7742
12.2674 | -104.7726
12.2711 | -104.777
12.2748 | -104.7828
12.2785 | -104.789
12.2822 | -104.7978
12.2858 | -104.8104
12.2895 | -104.8231
12.2932 | -104.8391
12.2969 | -104.8623
12.3006 | -104.8842
12.3043 | -104.9072
12.308 | -104.932
12.3117 | -104.9549
12.3154 | -104.9793
12.319 | -105.0033
12.3227 | -105.028
12.3264 | -105.0532
12.3301 | -105.0759
12.3338 | -105.0952
12.3375 | -105.1179
12.3412 | -105.1417
12.3449 | -105.1603
12.3485 | -105.1791
12.3522 | -105.1945
12.3559 | -105.2078
12.3596 | -105.2183
12.3633 | -105.2266
12.367 | -105.2346
12.3707 | -105.24
12.3744 | -105.2368
12.3781 | -105.2315
12.3817 | -105.2274
12.3854 | -105.2202
12.3891 | -105.2131
12.3928 | -105.2085
12.3965 | -105.2
12.4002 | -105.1903
12.4039 | -105.1818
12.4076 | -105.1717
12.4112 | -105.1629
12.4149 | -105.1551
12.4186 | -105.1448
12.4223 | -105.1363
12.426 | -105.1314
12.4297 | -105.1297
12.4334 | -105.1249
12.4371 | -105.1184
12.4408 | -105.1165
12.4444 | -105.1127
12.4481 | -105.1059
12.4518 | -105.1011
12.4555 | -105.095
12.4592 | -105.0889
12.4629 | -105.0822
12.4666 | -105.0709
12.4703 | -105.062
12.474 | -105.0528
12.4776 | -105.0429
12.4813 | -105.0337
12.485 | -105.0215
12.4887 | -105.0103
12.4924 | -105.0003
12.4961 | -104.9875
12.4998 | -104.9738
12.5035 | -104.9621
12.5071 | -104.9508
12.5108 | -104.9385
12.5145 | -104.9291
12.5182 | -104.9176
12.5219 | -104.9049
12.5256 | -104.8965
12.5293 | -104.8898
12.533 | -104.8844
12.5367 | -104.882
12.5403 | -104.8838
12.544 | -104.8907
12.5477 | -104.9003
12.5514 | -104.9102
12.5551 | -104.9246
12.5588 | -104.9387
12.5625 | -104.9556
12.5662 | -104.9761
12.5698 | -104.9998
12.5735 | -105.0233
12.5772 | -105.0446
12.5809 | -105.0687
12.5846 | -105.0911
12.5883 | -105.1158
12.592 | -105.142
12.5957 | -105.1655
12.5994 | -105.187
12.603 | -105.206
12.6067 | -105.2176
12.6104 | -105.2289
12.6141 | -105.2398
12.6178 | -105.2475
12.6215 | -105.2519
12.6252 | -105.2577
12.6289 | -105.261
12.6326 | -105.2612
12.6362 | -105.2598
12.6399 | -105.2562
12.6436 | -105.2507
12.6473 | -105.2467
12.651 | -105.2418
12.6547 | -105.237
12.6584 | -105.2296
12.6621 | -105.2229
12.6657 | -105.2189
12.6694 | -105.2137
12.6731 | -105.2073
12.6768 | -105.2042
12.6805 | -105.2016
12.6842 | -105.1979
12.6879 | -105.1966
12.6916 | -105.1973
12.6953 | -105.1983
12.6989 | -105.1977
12.7026 | -105.199
12.7063 | -105.2003
12.71 | -105.2023
12.7137 | -105.2069
12.7174 | -105.2109
12.7211 | -105.2149
12.7248 | -105.2194
12.7284 | -105.2252
12.7321 | -105.2347
12.7358 | -105.2445
12.7395 | -105.2579
12.7432 | -105.2671
12.7469 | -105.2778
12.7506 | -105.2898
12.7543 | -105.2991
12.758 | -105.31
12.7616 | -105.321
12.7653 | -105.3325
12.769 | -105.3459
12.7727 | -105.3553
12.7764 | -105.3626
12.7801 | -105.3708
12.7838 | -105.3788
12.7875 | -105.3836
12.7911 | -105.3924
12.7948 | -105.3993
12.7985 | -105.4045
12.8022 | -105.4127
12.8059 | -105.4185
12.8096 | -105.4206
12.8133 | -105.4216
12.817 | -105.4244
12.8207 | -105.4278
12.8243 | -105.4257
12.828 | -105.4177
12.8317 | -105.4115
12.8354 | -105.4031
12.8391 | -105.3924
12.8428 | -105.3828
12.8465 | -105.3722
12.8502 | -105.3604
12.8538 | -105.3522
12.8575 | -105.3457
12.8612 | -105.3368
12.8649 | -105.3281
12.8686 | -105.3212
12.8723 | -105.3164
12.876 | -105.3121
12.8797 | -105.3105
12.8834 | -105.3123
12.887 | -105.3134
12.8907 | -105.3134
12.8944 | -105.3184
12.8981 | -105.3253
12.9018 | -105.334
12.9055 | -105.3417
12.9092 | -105.3478
12.9129 | -105.3555
12.9166 | -105.3659
12.9202 | -105.3761
12.9239 | -105.3873
12.9276 | -105.4028
12.9313 | -105.415
12.935 | -105.4262
12.9387 | -105.4419
12.9424 | -105.4557
12.9461 | -105.4657
12.9497 | -105.4781
12.9534 | -105.4903
12.9571 | -105.5016
12.9608 | -105.5136
12.9645 | -105.5248
12.9682 | -105.5349
12.9719 | -105.5482
12.9756 | -105.5615
12.9793 | -105.5666
12.9829 | -105.5678
12.9866 | -105.5706
12.9903 | -105.5743
12.994 | -105.5769
12.9977 | -105.5797
13.0014 | -105.5825
13.0051 | -105.5829
13.0088 | -105.5835
13.0124 | -105.5876
13.0161 | -105.5912
13.0198 | -105.5876
13.0235 | -105.5827
13.0272 | -105.5781
13.0309 | -105.5729
13.0346 | -105.5678
13.0383 | -105.5658
13.042 | -105.5625
13.0456 | -105.5578
13.0493 | -105.5547
13.053 | -105.5517
13.0567 | -105.5451
13.0604 | -105.5371
13.0641 | -105.5319
13.0678 | -105.5254
13.0715 | -105.5166
13.0752 | -105.5097
13.0788 | -105.5044
13.0825 | -105.4963
13.0862 | -105.4877
13.0899 | -105.4771
13.0936 | -105.4674
13.0973 | -105.4603
13.101 | -105.4511
13.1047 | -105.4391
13.1083 | -105.4298
13.112 | -105.4201
13.1157 | -105.4082
13.1194 | -105.3999
13.1231 | -105.3953
13.1268 | -105.3919
13.1305 | -105.3888
13.1342 | -105.3842
13.1379 | -105.3775
13.1415 | -105.3717
13.1452 | -105.3684
13.1489 | -105.3616
13.1526 | -105.3547
13.1563 | -105.3497
13.16 | -105.3426
13.1637 | -105.3355
13.1674 | -105.3314
13.171 | -105.33
13.1747 | -105.3275
13.1784 | -105.3211
13.1821 | -105.3157
13.1858 | -105.3099
13.1895 | -105.3025
13.1932 | -105.2974
13.1969 | -105.2933
13.2006 | -105.2873
13.2042 | -105.2798
13.2079 | -105.2724
13.2116 | -105.2626
13.2153 | -105.2533
13.219 | -105.2434
13.2227 | -105.2297
13.2264 | -105.2131
13.2301 | -105.1983
13.2337 | -105.1848
13.2374 | -105.1721
13.2411 | -105.1604
13.2448 | -105.1508
13.2485 | -105.1388
13.2522 | -105.1264
13.2559 | -105.1151
13.2596 | -105.107
13.2633 | -105.0993
13.2669 | -105.0889
13.2706 | -105.0824
13.2743 | -105.0784
13.278 | -105.0738
13.2817 | -105.0698
13.2854 | -105.0658
13.2891 | -105.061
13.2928 | -105.0591
13.2965 | -105.057
13.3001 | -105.0583
13.3038 | -105.0571
13.3075 | -105.0568
13.3112 | -105.0581
13.3149 | -105.0584
13.3186 | -105.0587
13.3223 | -105.0597
13.326 | -105.0639
13.3296 | -105.0708
13.3333 | -105.0746
13.337 | -105.0741
13.3407 | -105.0763
13.3444 | -105.0805
13.3481 | -105.0825
13.3518 | -105.0819
13.3555 | -105.083
13.3592 | -105.0836
13.3628 | -105.0818
13.3665 | -105.0792
13.3702 | -105.0731
13.3739 | -105.0669
13.3776 | -105.0634
13.3813 | -105.0622
13.385 | -105.057
13.3887 | -105.0509
13.3923 | -105.0463
13.396 | -105.0391
13.3997 | -105.03
13.4034 | -105.0197
13.4071 | -105.0058
13.4108 | -104.9929
13.4145 | -104.9821
13.4182 | -104.971
13.4219 | -104.9584
13.4255 | -104.9472
13.4292 | -104.9371
13.4329 | -104.9275
13.4366 | -104.9206
13.4403 | -104.9142
13.444 | -104.9102
13.4477 | -104.9066
13.4514 | -104.9045
13.455 | -104.9039
13.4587 | -104.9097
13.4624 | -104.9159
13.4661 | -104.9228
13.4698 | -104.9338
13.4735 | -104.9448
13.4772 | -104.9587
13.4809 | -104.9761
13.4846 | -104.9946
13.4882 | -105.0135
13.4919 | -105.033
13.4956 | -105.055
13.4993 | -105.0767
13.503 | -105.1002
13.5067 | -105.1223
13.5104 | -105.144
13.5141 | -105.1647
13.5177 | -105.1847
13.5214 | -105.2018
13.5251 | -105.2173
13.5288 | -105.233
13.5325 | -105.2493
13.5362 | -105.2638
13.5399 | -105.2716
13.5436 | -105.2784
13.5473 | -105.2849
13.5509 | -105.288
13.5546 | -105.2879
13.5583 | -105.2852
13.562 | -105.283
13.5657 | -105.2802
13.5694 | -105.2729
13.5731 | -105.2625
13.5768 | -105.2521
13.5805 | -105.2448
13.5841 | -105.2346
13.5878 | -105.2232
13.5915 | -105.214
13.5952 | -105.2073
13.5989 | -105.2033
13.6026 | -105.199
13.6063 | -105.1913
13.61 | -105.1857
13.6136 | -105.1825
13.6173 | -105.1791
13.621 | -105.1743
13.6247 | -105.1684
13.6284 | -105.1653
13.6321 | -105.1617
13.6358 | -105.1591
13.6395 | -105.1579
13.6432 | -105.1596
13.6468 | -105.1605
13.6505 | -105.1615
13.6542 | -105.162
13.6579 | -105.1614
13.6616 | -105.1617
13.6653 | -105.1619
13.669 | -105.1621
13.6727 | -105.1571
13.6763 | -105.1524
13.68 | -105.1475
13.6837 | -105.1423
13.6874 | -105.1403
13.6911 | -105.1381
13.6948 | -105.1373
13.6985 | -105.1359
13.7022 | -105.1381
13.7059 | -105.1389
13.7095 | -105.1395
13.7132 | -105.1392
13.7169 | -105.1428
13.7206 | -105.1496
13.7243 | -105.1554
13.728 | -105.1609
13.7317 | -105.1655
13.7354 | -105.1712
13.7391 | -105.1775
13.7427 | -105.1841
13.7464 | -105.1912
13.7501 | -105.1978
13.7538 | -105.2066
13.7575 | -105.2149
13.7612 | -105.2214
13.7649 | -105.2303
13.7686 | -105.2385
13.7722 | -105.2425
13.7759 | -105.2432
13.7796 | -105.2459
13.7833 | -105.2511
13.787 | -105.2562
13.7907 | -105.2618
13.7944 | -105.2639
13.7981 | -105.2674
13.8018 | -105.2736
13.8054 | -105.2743
13.8091 | -105.2776
13.8128 | -105.2842
13.8165 | -105.292
13.8202 | -105.2994
13.8239 | -105.3074
13.8276 | -105.3174
13.8313 | -105.3295
13.8349 | -105.3395
13.8386 | -105.3495
13.8423 | -105.3589
13.846 | -105.3669
13.8497 | -105.3763
13.8534 | -105.3867
13.8571 | -105.4003
13.8608 | -105.4145
13.8645 | -105.4282
13.8681 | -105.4391
13.8718 | -105.4505
13.8755 | -105.4627
13.8792 | -105.4742
13.8829 | -105.4842
13.8866 | -105.4896
13.8903 | -105.4914
13.894 | -105.4936
13.8976 | -105.4988
13.9013 | -105.5006
13.905 | -105.503
13.9087 | -105.5047
13.9124 | -105.5061
13.9161 | -105.5033
13.9198 | -105.4952
13.9235 | -105.4884
13.9272 | -105.4816
13.9308 | -105.4743
13.9345 | -105.4674
13.9382 | -105.4568
13.9419 | -105.4456
13.9456 | -105.4317
13.9493 | -105.4214
13.953 | -105.4102
13.9567 | -105.3975
13.9604 | -105.3792
13.964 | -105.3591
13.9677 | -105.3427
13.9714 | -105.3304
13.9751 | -105.3201
13.9788 | -105.3114
13.9825 | -105.2986
13.9862 | -105.2856
13.9899 | -105.2734
13.9935 | -105.2627
13.9972 | -105.2518
14.0009 | -105.2401
14.0046 | -105.2291
14.0083 | -105.212
14.012 | -105.1923
14.0157 | -105.1737
14.0194 | -105.1545
14.0231 | -105.1325
14.0267 | -105.1107
14.0304 | -105.0911
14.0341 | -105.0747
14.0378 | -105.0592
14.0415 | -105.042
14.0452 | -105.0237
14.0489 | -105.0105
14.0526 | -105
14.0562 | -104.9906
14.0599 | -104.9829
14.0636 | -104.9785
14.0673 | -104.9797
14.071 | -104.989
14.0747 | -105.0026
14.0784 | -105.0154
14.0821 | -105.0302
14.0858 | -105.0491
14.0894 | -105.071
14.0931 | -105.0957
14.0968 | -105.1233
14.1005 | -105.1498
14.1042 | -105.1773
14.1079 | -105.2086
14.1116 | -105.2408
14.1153 | -105.2722
14.1189 | -105.3021
14.1226 | -105.3307
14.1263 | -105.3628
14.13 | -105.3948
14.1337 | -105.4232
14.1374 | -105.4498
14.1411 | -105.4764
14.1448 | -105.4994
14.1485 | -105.5232
14.1521 | -105.5449
14.1558 | -105.5636
14.1595 | -105.5797
14.1632 | -105.5998
14.1669 | -105.6165
14.1706 | -105.6244
14.1743 | -105.6274
14.178 | -105.63
14.1817 | -105.6289
14.1853 | -105.6293
14.189 | -105.6244
14.1927 | -105.6169
14.1964 | -105.6096
14.2001 | -105.6016
14.2038 | -105.5939
14.2075 | -105.5871
14.2112 | -105.5815
14.2148 | -105.5775
14.2185 | -105.5726
14.2222 | -105.5673
14.2259 | -105.5604
14.2296 | -105.5508
14.2333 | -105.5411
14.237 | -105.5351
14.2407 | -105.5309
14.2444 | -105.5278
14.248 | -105.5273
14.2517 | -105.5279
14.2554 | -105.5274
14.2591 | -105.5269
14.2628 | -105.5283
14.2665 | -105.5323
14.2702 | -105.5322
14.2739 | -105.5318
14.2775 | -105.5309
14.2812 | -105.5301
14.2849 | -105.53
14.2886 | -105.5299
14.2923 | -105.5303
14.296 | -105.529
14.2997 | -105.5268
14.3034 | -105.5232
14.3071 | -105.5177
14.3107 | -105.5077
14.3144 | -105.4978
14.3181 | -105.4876
14.3218 | -105.477
14.3255 | -105.4633
14.3292 | -105.442
14.3329 | -105.4203
14.3366 | -105.402
14.3402 | -105.3841
14.3439 | -105.365
14.3476 | -105.348
14.3513 | -105.3285
14.355 | -105.3064
14.3587 | -105.2866
14.3624 | -105.2689
14.3661 | -105.259
14.3698 | -105.2446
14.3734 | -105.2305
14.3771 | -105.22
14.3808 | -105.211
14.3845 | -105.2009
14.3882 | -105.1933
14.3919 | -105.1894
14.3956 | -105.1916
14.3993 | -105.1952
14.403 | -105.2003
14.4066 | -105.2068
14.4103 | -105.2113
14.414 | -105.2171
14.4177 | -105.2258
14.4214 | -105.2339
14.4251 | -105.2424
14.4288 | -105.2557
14.4325 | -105.2732
14.4361 | -105.2908
14.4398 | -105.3064
14.4435 | -105.3183
14.4472 | -105.3308
14.4509 | -105.3455
14.4546 | -105.3579
14.4583 | -105.3678
14.462 | -105.3786
14.4657 | -105.3897
14.4693 | -105.3979
14.473 | -105.4027
14.4767 | -105.4115
14.4804 | -105.4194
14.4841 | -105.4255
14.4878 | -105.4305
14.4915 | -105.4362
14.4952 | -105.4402
14.4988 | -105.4418
14.5025 | -105.4425
14.5062 | -105.4417
14.5099 | -105.4423
14.5136 | -105.4462
14.5173 | -105.4473
14.521 | -105.4469
14.5247 | -105.4478
14.5284 | -105.4505
14.532 | -105.4558
14.5357 | -105.4635
14.5394 | -105.4674
14.5431 | -105.4683
14.5468 | -105.4694
14.5505 | -105.4705
14.5542 | -105.4675
14.5579 | -105.4663
14.5615 | -105.4649
14.5652 | -105.4621
14.5689 | -105.4618
14.5726 | -105.4608
14.5763 | -105.4603
14.58 | -105.4586
14.5837 | -105.4548
14.5874 | -105.4488
14.5911 | -105.4454
14.5947 | -105.443
14.5984 | -105.4418
14.6021 | -105.4387
14.6058 | -105.4371
14.6095 | -105.4339
14.6132 | -105.4328
14.6169 | -105.4316
14.6206 | -105.4352
14.6243 | -105.4398
14.6279 | -105.4424
14.6316 | -105.4458
14.6353 | -105.4522
14.639 | -105.4558
14.6427 | -105.4597
14.6464 | -105.4692
14.6501 | -105.4818
14.6538 | -105.4946
14.6574 | -105.5071
14.6611 | -105.5194
14.6648 | -105.5356
14.6685 | -105.5517
14.6722 | -105.5647
14.6759 | -105.5783
14.6796 | -105.5945
14.6833 | -105.6148
14.687 | -105.6305
14.6906 | -105.6443
14.6943 | -105.6595
14.698 | -105.6763
14.7017 | -105.6919
14.7054 | -105.7079
14.7091 | -105.7253
14.7128 | -105.7414
14.7165 | -105.7532
14.7201 | -105.7642
14.7238 | -105.7749
14.7275 | -105.7879
14.7312 | -105.7976
14.7349 | -105.8056
14.7386 | -105.8105
14.7423 | -105.8164
14.746 | -105.8228
14.7497 | -105.8292
14.7533 | -105.8336
14.757 | -105.8361
14.7607 | -105.8383
14.7644 | -105.8414
14.7681 | -105.8414
14.7718 | -105.8416
14.7755 | -105.8383
14.7792 | -105.8315
14.7828 | -105.8251
14.7865 | -105.8204
14.7902 | -105.8153
14.7939 | -105.8108
14.7976 | -105.8071
14.8013 | -105.7999
14.805 | -105.7912
14.8087 | -105.782
14.8124 | -105.7735
14.816 | -105.7635
14.8197 | -105.7555
14.8234 | -105.7465
14.8271 | -105.7383
14.8308 | -105.7281
14.8345 | -105.7193
14.8382 | -105.7117
14.8419 | -105.7026
14.8456 | -105.6932
14.8492 | -105.6872
14.8529 | -105.6817
14.8566 | -105.6693
14.8603 | -105.6562
14.864 | -105.6472
14.8677 | -105.6393
14.8714 | -105.6344
14.8751 | -105.6283
14.8787 | -105.6179
14.8824 | -105.606
14.8861 | -105.5944
14.8898 | -105.5843
14.8935 | -105.5733
14.8972 | -105.5618
14.9009 | -105.5514
14.9046 | -105.5395
14.9083 | -105.525
14.9119 | -105.5108
14.9156 | -105.4974
14.9193 | -105.4845
14.923 | -105.4722
14.9267 | -105.461
14.9304 | -105.4532
14.9341 | -105.4426
14.9378 | -105.4286
14.9414 | -105.4153
14.9451 | -105.4002
14.9488 | -105.3888
14.9525 | -105.3778
14.9562 | -105.3669
14.9599 | -105.3561
14.9636 | -105.3474
14.9673 | -105.34
14.971 | -105.3342
14.9746 | -105.3298
14.9783 | -105.3235
14.982 | -105.3175
14.9857 | -105.3112
14.9894 | -105.3071
14.9931 | -105.3044
14.9968 | -105.303
15.0005 | -105.3016
15.0041 | -105.3021
15.0078 | -105.3016
15.0115 | -105.3017
15.0152 | -105.3015
15.0189 | -105.2994
15.0226 | -105.2996
15.0263 | -105.2979
15.03 | -105.2959
15.0337 | -105.2941
15.0373 | -105.2927
15.041 | -105.2914
15.0447 | -105.2918
15.0484 | -105.2889
15.0521 | -105.2869
15.0558 | -105.2883
15.0595 | -105.2887
15.0632 | -105.2874
15.0669 | -105.2849
15.0705 | -105.2857
15.0742 | -105.2902
15.0779 | -105.2924
15.0816 | -105.2943
15.0853 | -105.2958
15.089 | -105.3022
15.0927 | -105.3124
15.0964 | -105.3259
15.1 | -105.3361
15.1037 | -105.3444
15.1074 | -105.3532
15.1111 | -105.3619
15.1148 | -105.3695
15.1185 | -105.379
15.1222 | -105.3917
15.1259 | -105.4063
15.1296 | -105.4203
15.1332 | -105.4341
15.1369 | -105.4491
15.1406 | -105.4616
15.1443 | -105.4741
15.148 | -105.4878
15.1517 | -105.5024
15.1554 | -105.5187
15.1591 | -105.5327
15.1627 | -105.5479
15.1664 | -105.5619
15.1701 | -105.5761
15.1738 | -105.5899
15.1775 | -105.604
15.1812 | -105.6176
15.1849 | -105.6306
15.1886 | -105.6436
15.1923 | -105.6563
15.1959 | -105.6707
15.1996 | -105.6838
15.2033 | -105.6964
15.207 | -105.7084
15.2107 | -105.7183
15.2144 | -105.728
15.2181 | -105.7356
15.2218 | -105.7444
15.2255 | -105.755
15.2291 | -105.765
15.2328 | -105.7739
15.2365 | -105.7857
15.2402 | -105.7944
15.2439 | -105.8026
15.2476 | -105.8126
15.2513 | -105.8232
15.255 | -105.8301
15.2586 | -105.835
15.2623 | -105.8377
15.266 | -105.8433
15.2697 | -105.8457
15.2734 | -105.8442
15.2771 | -105.8405
15.2808 | -105.8349
15.2845 | -105.8306
15.2882 | -105.8295
15.2918 | -105.8216
15.2955 | -105.812
15.2992 | -105.8036
15.3029 | -105.7931
15.3066 | -105.7787
15.3103 | -105.7606
15.314 | -105.7397
15.3177 | -105.7211
15.3213 | -105.6995
15.325 | -105.6753
15.3287 | -105.6534
15.3324 | -105.6329
15.3361 | -105.6102
15.3398 | -105.5874
15.3435 | -105.5656
15.3472 | -105.5443
15.3509 | -105.5213
15.3545 | -105.4978
15.3582 | -105.4773
15.3619 | -105.4596
15.3656 | -105.4449
15.3693 | -105.4309
15.373 | -105.4181
15.3767 | -105.4087
15.3804 | -105.4
15.384 | -105.3908
15.3877 | -105.3875
15.3914 | -105.3868
15.3951 | -105.3849
15.3988 | -105.3867
15.4025 | -105.3925
15.4062 | -105.4007
15.4099 | -105.4079
15.4136 | -105.4176
15.4172 | -105.4312
15.4209 | -105.4453
15.4246 | -105.4578
15.4283 | -105.4738
15.432 | -105.4868
15.4357 | -105.499
15.4394 | -105.5134
15.4431 | -105.5258
15.4467 | -105.5376
15.4504 | -105.5492
15.4541 | -105.562
15.4578 | -105.5745
15.4615 | -105.5845
15.4652 | -105.5942
15.4689 | -105.6021
15.4726 | -105.6078
15.4763 | -105.615
15.4799 | -105.6223
15.4836 | -105.6275
15.4873 | -105.6295
15.491 | -105.6327
15.4947 | -105.6345
15.4984 | -105.6356
15.5021 | -105.6383
15.5058 | -105.6377
15.5095 | -105.6335
15.5131 | -105.6323
15.5168 | -105.6323
15.5205 | -105.6312
15.5242 | -105.6284
15.5279 | -105.6259
15.5316 | -105.6267
15.5353 | -105.6275
15.539 | -105.6274
15.5426 | -105.6272
15.5463 | -105.6277
15.55 | -105.6261
15.5537 | -105.6247
15.5574 | -105.6259
15.5611 | -105.6257
15.5648 | -105.6249
15.5685 | -105.6231
15.5722 | -105.6252
15.5758 | -105.6265
15.5795 | -105.6296
15.5832 | -105.6301
15.5869 | -105.6254
15.5906 | -105.6232
15.5943 | -105.6208
15.598 | -105.6162
15.6017 | -105.6112
15.6053 | -105.6035
15.609 | -105.5911
15.6127 | -105.5776
15.6164 | -105.5647
15.6201 | -105.5494
15.6238 | -105.5325
15.6275 | -105.5126
15.6312 | -105.4939
15.6349 | -105.4748
15.6385 | -105.4552
15.6422 | -105.4343
15.6459 | -105.4163
15.6496 | -105.3979
15.6533 | -105.3776
15.657 | -105.3577
15.6607 | -105.3389
15.6644 | -105.3207
15.668 | -105.3023
15.6717 | -105.2855
15.6754 | -105.2714
15.6791 | -105.258
15.6828 | -105.2432
15.6865 | -105.2313
15.6902 | -105.2229
15.6939 | -105.2155
15.6976 | -105.2094
15.7012 | -105.2037
15.7049 | -105.2001
15.7086 | -105.1987
15.7123 | -105.2003
15.716 | -105.2033
15.7197 | -105.2087
15.7234 | -105.2145
15.7271 | -105.2192
15.7308 | -105.2269
15.7344 | -105.2379
15.7381 | -105.2505
15.7418 | -105.263
15.7455 | -105.2755
15.7492 | -105.2902
15.7529 | -105.3068
15.7566 | -105.3231
15.7603 | -105.3386
15.7639 | -105.3532
15.7676 | -105.3679
15.7713 | -105.3806
15.775 | -105.3942
15.7787 | -105.4034
15.7824 | -105.4102
15.7861 | -105.4156
15.7898 | -105.4191
15.7935 | -105.4226
15.7971 | -105.425
15.8008 | -105.4285
15.8045 | -105.4297
15.8082 | -105.4301
15.8119 | -105.4296
15.8156 | -105.4253
15.8193 | -105.422
15.823 | -105.4198
15.8266 | -105.4145
15.8303 | -105.4076
15.834 | -105.3995
15.8377 | -105.389
15.8414 | -105.3816
15.8451 | -105.3747
15.8488 | -105.3659
15.8525 | -105.355
15.8562 | -105.3448
15.8598 | -105.3353
15.8635 | -105.3226
15.8672 | -105.3108
15.8709 | -105.3
15.8746 | -105.2891
15.8783 | -105.2798
15.882 | -105.2698
15.8857 | -105.2576
15.8894 | -105.2458
15.893 | -105.2364
15.8967 | -105.2275
15.9004 | -105.2168
15.9041 | -105.2108
15.9078 | -105.2075
15.9115 | -105.2045
15.9152 | -105.2014
15.9189 | -105.1996
15.9225 | -105.1998
15.9262 | -105.2038
15.9299 | -105.2074
15.9336 | -105.2118
15.9373 | -105.2174
15.941 | -105.2239
15.9447 | -105.2333
15.9484 | -105.2452
15.9521 | -105.2585
15.9557 | -105.2712
15.9594 | -105.2828
15.9631 | -105.3009
15.9668 | -105.3115
15.9705 | -105.3243
15.9742 | -105.3376
15.9779 | -105.3517
15.9816 | -105.3667
15.9852 | -105.3833
15.9889 | -105.3994
15.9926 | -105.4134
15.9963 | -105.4283
16 | -105.4434
16.0037 | -105.4561
16.0074 | -105.4691
16.0111 | -105.4809
16.0148 | -105.4903
16.0184 | -105.5008
16.0221 | -105.5092
16.0258 | -105.5122
16.0295 | -105.5135
16.0332 | -105.5177
16.0369 | -105.5215
16.0406 | -105.5273
16.0443 | -105.5331
16.0479 | -105.5329
16.0516 | -105.5303
16.0553 | -105.5281
16.059 | -105.5245
16.0627 | -105.5223
16.0664 | -105.5225
16.0701 | -105.5194
16.0738 | -105.5118
16.0775 | -105.504
16.0811 | -105.4986
16.0848 | -105.4924
16.0885 | -105.4839
16.0922 | -105.4775
16.0959 | -105.4712
16.0996 | -105.4631
16.1033 | -105.4559
16.107 | -105.4481
16.1106 | -105.44
16.1143 | -105.4339
16.118 | -105.4267
16.1217 | -105.4194
16.1254 | -105.4124
16.1291 | -105.4058
16.1328 | -105.3983
16.1365 | -105.3892
16.1402 | -105.3807
16.1438 | -105.3729
16.1475 | -105.3658
16.1512 | -105.3601
16.1549 | -105.3539
16.1586 | -105.3508
16.1623 | -105.3497
16.166 | -105.347
16.1697 | -105.3441
16.1734 | -105.3435
16.177 | -105.344
16.1807 | -105.3432
16.1844 | -105.342
16.1881 | -105.3429
16.1918 | -105.3435
16.1955 | -105.346
16.1992 | -105.3477
16.2029 | -105.3501
16.2065 | -105.3551
16.2102 | -105.3596
16.2139 | -105.3611
16.2176 | -105.3626
16.2213 | -105.366
16.225 | -105.3698
16.2287 | -105.372
16.2324 | -105.373
16.2361 | -105.3746
16.2397 | -105.3761
16.2434 | -105.3766
16.2471 | -105.3778
16.2508 | -105.3774
16.2545 | -105.3779
16.2582 | -105.3797
16.2619 | -105.3816
16.2656 | -105.3811
16.2692 | -105.3822
16.2729 | -105.3828
16.2766 | -105.3825
16.2803 | -105.3831
16.284 | -105.3857
16.2877 | -105.3911
16.2914 | -105.3939
16.2951 | -105.3938
16.2988 | -105.3926
16.3024 | -105.3911
16.3061 | -105.3946
16.3098 | -105.4001
16.3135 | -105.4043
16.3172 | -105.4068
16.3209 | -105.4083
16.3246 | -105.4107
16.3283 | -105.4137
16.3319 | -105.4165
16.3356 | -105.4199
16.3393 | -105.4256
16.343 | -105.4337
16.3467 | -105.4403
16.3504 | -105.4481
16.3541 | -105.4553
16.3578 | -105.4618
16.3615 | -105.4696
16.3651 | -105.4763
16.3688 | -105.4862
16.3725 | -105.4982
16.3762 | -105.5109
16.3799 | -105.5235
16.3836 | -105.5347
16.3873 | -105.5452
16.391 | -105.5539
16.3947 | -105.5604
16.3983 | -105.5695
16.402 | -105.5822
16.4057 | -105.5923
16.4094 | -105.6012
16.4131 | -105.6119
16.4168 | -105.6226
16.4205 | -105.6313
16.4242 | -105.6398
16.4278 | -105.6473
16.4315 | -105.6554
16.4352 | -105.6617
16.4389 | -105.6667
16.4426 | -105.6699
16.4463 | -105.6727
16.45 | -105.6736
16.4537 | -105.6737
16.4574 | -105.6712
16.461 | -105.6688
16.4647 | -105.6677
16.4684 | -105.6643
16.4721 | -105.6581
16.4758 | -105.6496
16.4795 | -105.6416
16.4832 | -105.6325
16.4869 | -105.6188
16.4905 | -105.6073
16.4942 | -105.5992
16.4979 | -105.5892
16.5016 | -105.5794
16.5053 | -105.5684
16.509 | -105.5589
16.5127 | -105.5503
16.5164 | -105.5443
16.5201 | -105.5395
16.5237 | -105.5326
16.5274 | -105.5281
16.5311 | -105.5239
16.5348 | -105.5204
16.5385 | -105.5187
16.5422 | -105.515
16.5459 | -105.5113
16.5496 | -105.5091
16.5532 | -105.5082
16.5569 | -105.5111
16.5606 | -105.5135
16.5643 | -105.5163
16.568 | -105.5211
16.5717 | -105.5227
16.5754 | -105.5233
16.5791 | -105.5245
16.5828 | -105.5276
16.5864 | -105.5303
16.5901 | -105.5355
16.5938 | -105.5419
16.5975 | -105.5488
16.6012 | -105.558
16.6049 | -105.5674
16.6086 | -105.574
16.6123 | -105.5799
16.616 | -105.5881
16.6196 | -105.5979
16.6233 | -105.6086
16.627 | -105.618
16.6307 | -105.6291
16.6344 | -105.6371
16.6381 | -105.6429
16.6418 | -105.649
16.6455 | -105.6577
16.6491 | -105.6667
16.6528 | -105.6737
16.6565 | -105.6787
16.6602 | -105.6855
16.6639 | -105.6928
16.6676 | -105.6969
16.6713 | -105.6998
16.675 | -105.7022
16.6787 | -105.708
16.6823 | -105.7136
16.686 | -105.7178
16.6897 | -105.7223
16.6934 | -105.7293
16.6971 | -105.7351
16.7008 | -105.7377
16.7045 | -105.74
16.7082 | -105.744
16.7118 | -105.7488
16.7155 | -105.7536
16.7192 | -105.7566
16.7229 | -105.7582
16.7266 | -105.7589
16.7303 | -105.7578
16.734 | -105.7571
16.7377 | -105.7565
16.7414 | -105.7552
16.745 | -105.7566
16.7487 | -105.7583
16.7524 | -105.76
16.7561 | -105.7604
16.7598 | -105.7581
16.7635 | -105.7573
16.7672 | -105.7583
16.7709 | -105.7615
16.7746 | -105.7622
16.7782 | -105.7636
16.7819 | -105.7665
16.7856 | -105.7702
16.7893 | -105.7736
16.793 | -105.7776
16.7967 | -105.7776
16.8004 | -105.7777
16.8041 | -105.7779
16.8077 | -105.7789
16.8114 | -105.7844
16.8151 | -105.7889
16.8188 | -105.7919
16.8225 | -105.796
16.8262 | -105.7991
16.8299 | -105.7991
16.8336 | -105.8007
16.8373 | -105.8002
16.8409 | -105.8001
16.8446 | -105.8
16.8483 | -105.8037
16.852 | -105.807
16.8557 | -105.8098
16.8594 | -105.8138
16.8631 | -105.8184
16.8668 | -105.8235
16.8704 | -105.8288
16.8741 | -105.8361
16.8778 | -105.8417
16.8815 | -105.8475
16.8852 | -105.8517
16.8889 | -105.8548
16.8926 | -105.86
16.8963 | -105.867
16.9 | -105.8719
16.9036 | -105.8768
16.9073 | -105.8843
16.911 | -105.8907
16.9147 | -105.8958
16.9184 | -105.9005
16.9221 | -105.9048
16.9258 | -105.9103
16.9295 | -105.9151
16.9331 | -105.9196
16.9368 | -105.9249
16.9405 | -105.9312
16.9442 | -105.9393
16.9479 | -105.9463
16.9516 | -105.952
16.9553 | -105.9572
16.959 | -105.9634
16.9627 | -105.9728
16.9663 | -105.9795
16.97 | -105.9845
16.9737 | -105.9903
16.9774 | -105.9973
16.9811 | -106.0032
16.9848 | -106.007
16.9885 | -106.0108
16.9922 | -106.015
16.9959 | -106.0174
16.9995 | -106.0203
17.0032 | -106.0217
17.0069 | -106.0233
17.0106 | -106.0238
17.0143 | -106.0242
17.018 | -106.0203
17.0217 | -106.0159
17.0254 | -106.013
17.029 | -106.0091
17.0327 | -106.0051
17.0364 | -105.9994
17.0401 | -105.9967
17.0438 | -105.9943
17.0475 | -105.9911
17.0512 | -105.9861
17.0549 | -105.9789
17.0586 | -105.9746
17.0622 | -105.9687
17.0659 | -105.96
17.0696 | -105.9544
17.0733 | -105.9513
17.077 | -105.9475
17.0807 | -105.9459
17.0844 | -105.9404
17.0881 | -105.9344
17.0917 | -105.9286
17.0954 | -105.9228
17.0991 | -105.9161
17.1028 | -105.9103
17.1065 | -105.9059
17.1102 | -105.9014
17.1139 | -105.8967
17.1176 | -105.8927
17.1213 | -105.8879
17.1249 | -105.8843
17.1286 | -105.8824
17.1323 | -105.8807
17.136 | -105.8801
17.1397 | -105.879
17.1434 | -105.8764
17.1471 | -105.875
17.1508 | -105.8723
17.1544 | -105.8712
17.1581 | -105.8716
17.1618 | -105.8706
17.1655 | -105.8713
17.1692 | -105.8715
17.1729 | -105.8735
17.1766 | -105.8748
17.1803 | -105.8755
17.184 | -105.876
17.1876 | -105.877
17.1913 | -105.8785
17.195 | -105.8798
17.1987 | -105.8816
17.2024 | -105.8835
17.2061 | -105.8832
17.2098 | -105.8819
17.2135 | -105.882
17.2172 | -105.8808
17.2208 | -105.8777
17.2245 | -105.8749
17.2282 | -105.8701
17.2319 | -105.8648
17.2356 | -105.8597
17.2393 | -105.8549
17.243 | -105.8491
17.2467 | -105.8424
17.2503 | -105.8352
17.254 | -105.8262
17.2577 | -105.8194
17.2614 | -105.815
17.2651 | -105.8099
17.2688 | -105.8029
17.2725 | -105.7957
17.2762 | -105.7901
17.2799 | -105.784
17.2835 | -105.7804
17.2872 | -105.7735
17.2909 | -105.7664
17.2946 | -105.7597
17.2983 | -105.7548
17.302 | -105.7522
17.3057 | -105.7472
17.3094 | -105.7441
17.313 | -105.7418
17.3167 | -105.7386
17.3204 | -105.7356
17.3241 | -105.7304
17.3278 | -105.7277
17.3315 | -105.7278
17.3352 | -105.7285
17.3389 | -105.7281
17.3426 | -105.7291
17.3462 | -105.733
17.3499 | -105.7375
17.3536 | -105.7424
17.3573 | -105.7492
17.361 | -105.7566
17.3647 | -105.7633
17.3684 | -105.772
17.3721 | -105.7819
17.3757 | -105.7939
17.3794 | -105.8062
17.3831 | -105.8181
17.3868 | -105.832
17.3905 | -105.8495
17.3942 | -105.8638
17.3979 | -105.8794
17.4016 | -105.8948
17.4053 | -105.9103
17.4089 | -105.9262
17.4126 | -105.9382
17.4163 | -105.9482
17.42 | -105.9553
17.4237 | -105.9635
17.4274 | -105.9711
17.4311 | -105.9786
17.4348 | -105.9837
17.4385 | -105.986
17.4421 | -105.9898
17.4458 | -105.9906
17.4495 | -105.9925
17.4532 | -105.9935
17.4569 | -105.9908
17.4606 | -105.9862
17.4643 | -105.9822
17.468 | -105.9778
17.4716 | -105.9713
17.4753 | -105.9658
17.479 | -105.96
17.4827 | -105.9527
17.4864 | -105.9444
17.4901 | -105.9336
17.4938 | -105.9223
17.4975 | -105.9137
17.5012 | -105.9061
17.5048 | -105.8962
17.5085 | -105.8866
17.5122 | -105.8773
17.5159 | -105.8687
17.5196 | -105.8588
17.5233 | -105.8509
17.527 | -105.8472
17.5307 | -105.8426
17.5343 | -105.8391
17.538 | -105.8358
17.5417 | -105.8322
17.5454 | -105.8275
17.5491 | -105.823
17.5528 | -105.8193
17.5565 | -105.8148
17.5602 | -105.808
17.5639 | -105.8017
17.5675 | -105.7946
17.5712 | -105.7858
17.5749 | -105.7785
17.5786 | -105.7702
17.5823 | -105.7631
17.586 | -105.7547
17.5897 | -105.7465
17.5934 | -105.739
17.597 | -105.727
17.6007 | -105.7135
17.6044 | -105.7007
17.6081 | -105.6898
17.6118 | -105.678
17.6155 | -105.6641
17.6192 | -105.6511
17.6229 | -105.639
17.6266 | -105.6252
17.6302 | -105.6131
17.6339 | -105.6
17.6376 | -105.5873
17.6413 | -105.5721
17.645 | -105.5588
17.6487 | -105.5469
17.6524 | -105.5348
17.6561 | -105.5203
17.6598 | -105.5056
17.6634 | -105.4946
17.6671 | -105.4861
17.6708 | -105.4788
17.6745 | -105.4705
17.6782 | -105.4626
17.6819 | -105.4538
17.6856 | -105.4447
17.6893 | -105.4393
17.6929 | -105.4331
17.6966 | -105.4274
17.7003 | -105.4221
17.704 | -105.417
17.7077 | -105.4149
17.7114 | -105.4146
17.7151 | -105.4122
17.7188 | -105.4119
17.7225 | -105.4146
17.7261 | -105.4162
17.7298 | -105.4204
17.7335 | -105.4265
17.7372 | -105.4338
17.7409 | -105.4412
17.7446 | -105.449
17.7483 | -105.4535
17.752 | -105.4602
17.7556 | -105.4675
17.7593 | -105.4745
17.763 | -105.4827
17.7667 | -105.4929
17.7704 | -105.5046
17.7741 | -105.5162
17.7778 | -105.5262
17.7815 | -105.535
17.7852 | -105.5453
17.7888 | -105.5548
17.7925 | -105.5659
17.7962 | -105.5773
17.7999 | -105.5891
17.8036 | -105.6013
17.8073 | -105.6125
17.811 | -105.6224
17.8147 | -105.6343
17.8183 | -105.6481
17.822 | -105.659
17.8257 | -105.6698
17.8294 | -105.682
17.8331 | -105.6923
17.8368 | -105.7006
17.8405 | -105.71
17.8442 | -105.7198
17.8479 | -105.7273
17.8515 | -105.7362
17.8552 | -105.7458
17.8589 | -105.7542
17.8626 | -105.7613
17.8663 | -105.7698
17.87 | -105.7809
17.8737 | -105.7902
17.8774 | -105.7972
17.8811 | -105.8044
17.8847 | -105.8085
17.8884 | -105.8111
17.8921 | -105.8154
17.8958 | -105.823
17.8995 | -105.829
17.9032 | -105.8328
17.9069 | -105.8358
17.9106 | -105.8377
17.9142 | -105.8372
17.9179 | -105.8355
17.9216 | -105.8325
17.9253 | -105.8285
17.929 | -105.8261
17.9327 | -105.8218
17.9364 | -105.8189
17.9401 | -105.8152
17.9438 | -105.8087
17.9474 | -105.8006
17.9511 | -105.7905
17.9548 | -105.7809
17.9585 | -105.7729
17.9622 | -105.7643
17.9659 | -105.7537
17.9696 | -105.7435
17.9733 | -105.7348
17.9769 | -105.7253
17.9806 | -105.7151
17.9843 | -105.7022
17.988 | -105.6907
17.9917 | -105.6778
17.9954 | -105.6666
17.9991 | -105.6582
18.0028 | -105.6482
18.0065 | -105.6376
18.0101 | -105.6277
18.0138 | -105.6195
18.0175 | -105.6108
18.0212 | -105.6002
18.0249 | -105.5916
18.0286 | -105.5818
18.0323 | -105.5729
18.036 | -105.5679
18.0397 | -105.5639
18.0433 | -105.5615
18.047 | -105.5571
18.0507 | -105.5535
18.0544 | -105.5497
18.0581 | -105.5478
18.0618 | -105.5489
18.0655 | -105.5524
18.0692 | -105.5546
18.0728 | -105.5552
18.0765 | -105.5608
18.0802 | -105.5666
18.0839 | -105.5717
18.0876 | -105.5801
18.0913 | -105.5935
18.095 | -105.603
18.0987 | -105.6124
18.1024 | -105.6188
18.106 | -105.6278
18.1097 | -105.6403
18.1134 | -105.6507
18.1171 | -105.6624
18.1208 | -105.6746
18.1245 | -105.6846
18.1282 | -105.695
18.1319 | -105.706
18.1355 | -105.7165
18.1392 | -105.7281
18.1429 | -105.7413
18.1466 | -105.7568
18.1503 | -105.7658
18.154 | -105.7716
18.1577 | -105.7789
18.1614 | -105.7856
18.1651 | -105.7914
18.1687 | -105.7975
18.1724 | -105.8023
18.1761 | -105.8042
18.1798 | -105.8022
18.1835 | -105.7981
18.1872 | -105.7905
18.1909 | -105.7852
18.1946 | -105.7791
18.1982 | -105.7713
18.2019 | -105.7632
18.2056 | -105.754
18.2093 | -105.7457
18.213 | -105.7364
18.2167 | -105.7255
18.2204 | -105.7131
18.2241 | -105.7001
18.2278 | -105.688
18.2314 | -105.6745
18.2351 | -105.6599
18.2388 | -105.6481
18.2425 | -105.6379
18.2462 | -105.6243
18.2499 | -105.6127
18.2536 | -105.602
18.2573 | -105.5892
18.261 | -105.5789
18.2646 | -105.5703
18.2683 | -105.5622
18.272 | -105.5532
18.2757 | -105.5458
18.2794 | -105.5371
18.2831 | -105.5291
18.2868 | -105.5249
18.2905 | -105.5221
18.2941 | -105.5203
18.2978 | -105.52
18.3015 | -105.5202
18.3052 | -105.5238
18.3089 | -105.5284
18.3126 | -105.5324
18.3163 | -105.5342
18.32 | -105.5383
18.3237 | -105.5457
18.3273 | -105.5547
18.331 | -105.5649
18.3347 | -105.576
18.3384 | -105.5881
18.3421 | -105.599
18.3458 | -105.6093
18.3495 | -105.6212
18.3532 | -105.6331
18.3568 | -105.645
18.3605 | -105.6576
18.3642 | -105.674
18.3679 | -105.6893
18.3716 | -105.7021
18.3753 | -105.7156
18.379 | -105.7283
18.3827 | -105.7387
18.3864 | -105.749
18.39 | -105.7588
18.3937 | -105.7687
18.3974 | -105.7787
18.4011 | -105.786
18.4048 | -105.7927
18.4085 | -105.7984
18.4122 | -105.8052
18.4159 | -105.8118
18.4195 | -105.8165
18.4232 | -105.8205
18.4269 | -105.8242
18.4306 | -105.8261
18.4343 | -105.8291
18.438 | -105.8315
18.4417 | -105.8338
18.4454 | -105.8325
18.4491 | -105.8312
18.4527 | -105.8268
18.4564 | -105.8198
18.4601 | -105.8145
18.4638 | -105.8094
18.4675 | -105.8024
18.4712 | -105.7935
18.4749 | -105.7828
18.4786 | -105.7734
18.4823 | -105.7647
18.4859 | -105.7558
18.4896 | -105.7465
18.4933 | -105.7347
18.497 | -105.7238
18.5007 | -105.7143
18.5044 | -105.7022
18.5081 | -105.692
18.5118 | -105.6814
18.5154 | -105.671
18.5191 | -105.6612
18.5228 | -105.6507
18.5265 | -105.6413
18.5302 | -105.6325
18.5339 | -105.6225
18.5376 | -105.6146
18.5413 | -105.609
18.545 | -105.6023
18.5486 | -105.5962
18.5523 | -105.5895
18.556 | -105.5833
18.5597 | -105.5766
18.5634 | -105.5724
18.5671 | -105.5699
18.5708 | -105.5671
18.5745 | -105.5678
18.5781 | -105.5707
18.5818 | -105.5704
18.5855 | -105.5665
18.5892 | -105.5642
18.5929 | -105.5639
18.5966 | -105.5627
18.6003 | -105.5636
18.604 | -105.5652
18.6077 | -105.5674
18.6113 | -105.5701
18.615 | -105.571
18.6187 | -105.5753
18.6224 | -105.5816
18.6261 | -105.5851
18.6298 | -105.5876
18.6335 | -105.5905
18.6372 | -105.5916
18.6408 | -105.594
18.6445 | -105.5953
18.6482 | -105.5971
18.6519 | -105.601
18.6556 | -105.6036
18.6593 | -105.603
18.663 | -105.6011
18.6667 | -105.6013
18.6704 | -105.6014
18.674 | -105.6016
18.6777 | -105.6002
18.6814 | -105.598
18.6851 | -105.5961
18.6888 | -105.5923
18.6925 | -105.5882
18.6962 | -105.5863
18.6999 | -105.582
18.7035 | -105.5767
18.7072 | -105.5705
18.7109 | -105.5659
18.7146 | -105.5622
18.7183 | -105.5583
18.722 | -105.5537
18.7257 | -105.5489
18.7294 | -105.5431
18.7331 | -105.5388
18.7367 | -105.5354
18.7404 | -105.5307
18.7441 | -105.5246
18.7478 | -105.5182
18.7515 | -105.5128
18.7552 | -105.5066
18.7589 | -105.5032
18.7626 | -105.5001
18.7663 | -105.497
18.7699 | -105.494
18.7736 | -105.4925
18.7773 | -105.4919
18.781 | -105.4897
18.7847 | -105.4867
18.7884 | -105.4843
18.7921 | -105.4817
18.7958 | -105.4791
18.7994 | -105.4776
18.8031 | -105.4734
18.8068 | -105.4719
18.8105 | -105.4716
18.8142 | -105.4711
18.8179 | -105.4725
18.8216 | -105.4721
18.8253 | -105.4748
18.829 | -105.4779
18.8326 | -105.4796
18.8363 | -105.4813
18.84 | -105.4825
18.8437 | -105.4855
18.8474 | -105.4894
18.8511 | -105.4966
18.8548 | -105.503
18.8585 | -105.5073
18.8621 | -105.5133
18.8658 | -105.5177
18.8695 | -105.5242
18.8732 | -105.5302
18.8769 | -105.5352
18.8806 | -105.5418
18.8843 | -105.5447
18.888 | -105.5491
18.8917 | -105.5543
18.8953 | -105.5599
18.899 | -105.567
18.9027 | -105.5743
18.9064 | -105.5832
18.9101 | -105.5918
18.9138 | -105.6029
18.9175 | -105.6121
18.9212 | -105.6197
18.9248 | -105.6285
18.9285 | -105.6392
18.9322 | -105.6496
18.9359 | -105.6596
18.9396 | -105.6698
18.9433 | -105.6806
18.947 | -105.6939
18.9507 | -105.7056
18.9544 | -105.7147
18.958 | -105.7256
18.9617 | -105.7381
18.9654 | -105.7487
18.9691 | -105.7603
18.9728 | -105.7726
18.9765 | -105.7853
18.9802 | -105.7963
18.9839 | -105.8053
18.9876 | -105.8131
18.9912 | -105.8215
18.9949 | -105.8291
18.9986 | -105.837
19.0023 | -105.8441
19.006 | -105.8479
19.0097 | -105.8528
19.0134 | -105.8562
19.0171 | -105.8578
19.0207 | -105.8569
19.0244 | -105.8564
19.0281 | -105.8535
19.0318 | -105.8509
19.0355 | -105.8466
19.0392 | -105.8438
19.0429 | -105.8391
19.0466 | -105.8309
19.0503 | -105.823
19.0539 | -105.815
19.0576 | -105.8061
19.0613 | -105.7961
19.065 | -105.786
19.0687 | -105.7763
19.0724 | -105.7659
19.0761 | -105.7543
19.0798 | -105.7433
19.0834 | -105.7311
19.0871 | -105.7195
19.0908 | -105.7083
19.0945 | -105.6954
19.0982 | -105.6813
19.1019 | -105.6685
19.1056 | -105.6588
19.1093 | -105.6489
19.113 | -105.6375
19.1166 | -105.6288
19.1203 | -105.6188
19.124 | -105.6095
19.1277 | -105.602
19.1314 | -105.5937
19.1351 | -105.5873
19.1388 | -105.581
19.1425 | -105.577
19.1461 | -105.5755
19.1498 | -105.5727
19.1535 | -105.5691
19.1572 | -105.5655
19.1609 | -105.561
19.1646 | -105.5557
19.1683 | -105.5493
19.172 | -105.5402
19.1757 | -105.533
19.1793 | -105.524
19.183 | -105.5144
19.1867 | -105.506
19.1904 | -105.4998
19.1941 | -105.4917
19.1978 | -105.4818
19.2015 | -105.4705
19.2052 | -105.4599
19.2089 | -105.4487
19.2125 | -105.4356
19.2162 | -105.4207
19.2199 | -105.4071
19.2236 | -105.3944
19.2273 | -105.3819
19.231 | -105.367
19.2347 | -105.3489
19.2384 | -105.331
19.242 | -105.3132
19.2457 | -105.2946
19.2494 | -105.2726
19.2531 | -105.2506
19.2568 | -105.2292
19.2605 | -105.2091
19.2642 | -105.1891
19.2679 | -105.1676
19.2716 | -105.1462
19.2752 | -105.127
19.2789 | -105.1099
19.2826 | -105.0915
19.2863 | -105.0729
19.29 | -105.055
19.2937 | -105.0368
19.2974 | -105.0177
19.3011 | -104.9998
19.3047 | -104.9809
19.3084 | -104.9611
19.3121 | -104.9403
19.3158 | -104.9231
19.3195 | -104.9097
19.3232 | -104.896
19.3269 | -104.882
19.3306 | -104.8684
19.3343 | -104.8549
19.3379 | -104.8439
19.3416 | -104.8344
19.3453 | -104.8252
19.349 | -104.8177
19.3527 | -104.8102
19.3564 | -104.8042
19.3601 | -104.8002
19.3638 | -104.799
19.3675 | -104.7984
19.3711 | -104.8001
19.3748 | -104.803
19.3785 | -104.8063
19.3822 | -104.8106
19.3859 | -104.8175
19.3896 | -104.8239
19.3933 | -104.8304
19.397 | -104.837
19.4006 | -104.8459
19.4043 | -104.8571
19.408 | -104.8693
19.4117 | -104.8841
19.4154 | -104.899
19.4191 | -104.912
19.4228 | -104.9271
19.4265 | -104.9433
19.4302 | -104.9626
19.4338 | -104.9811
19.4375 | -104.9983
19.4412 | -105.0158
19.4449 | -105.0345
19.4486 | -105.0538
19.4523 | -105.0694
19.456 | -105.0879
19.4597 | -105.1036
19.4633 | -105.1148
19.467 | -105.1256
19.4707 | -105.136
19.4744 | -105.1444
19.4781 | -105.1529
19.4818 | -105.1586
19.4855 | -105.1623
19.4892 | -105.1678
19.4929 | -105.1714
19.4965 | -105.1699
19.5002 | -105.1675
19.5039 | -105.1641
19.5076 | -105.161
19.5113 | -105.1562
19.515 | -105.1488
19.5187 | -105.1392
19.5224 | -105.129
19.526 | -105.1193
19.5297 | -105.1073
19.5334 | -105.0918
19.5371 | -105.0764
19.5408 | -105.0606
19.5445 | -105.0454
19.5482 | -105.0296
19.5519 | -105.0115
19.5556 | -104.9907
19.5592 | -104.9715
19.5629 | -104.9546
19.5666 | -104.9381
19.5703 | -104.9243
19.574 | -104.9126
19.5777 | -104.8995
19.5814 | -104.8842
19.5851 | -104.8702
19.5888 | -104.8584
19.5924 | -104.847
19.5961 | -104.8368
19.5998 | -104.8298
19.6035 | -104.824
19.6072 | -104.8196
19.6109 | -104.8172
19.6146 | -104.8141
19.6183 | -104.8115
19.6219 | -104.8109
19.6256 | -104.8103
19.6293 | -104.8118
19.633 | -104.8149
19.6367 | -104.8172
19.6404 | -104.8198
19.6441 | -104.822
19.6478 | -104.8241
19.6515 | -104.8287
19.6551 | -104.8339
19.6588 | -104.8378
19.6625 | -104.8402
19.6662 | -104.8438
19.6699 | -104.8488
19.6736 | -104.8556
19.6773 | -104.86
19.681 | -104.8632
19.6846 | -104.8629
19.6883 | -104.8633
19.692 | -104.8627
19.6957 | -104.8619
19.6994 | -104.8638
19.7031 | -104.8647
19.7068 | -104.8652
19.7105 | -104.8651
19.7142 | -104.8642
19.7178 | -104.8637
19.7215 | -104.8645
19.7252 | -104.866
19.7289 | -104.8651
19.7326 | -104.8652
19.7363 | -104.8638
19.74 | -104.8649
19.7437 | -104.8675
19.7473 | -104.8689
19.751 | -104.8718
19.7547 | -104.8758
19.7584 | -104.8789
19.7621 | -104.8833
19.7658 | -104.8886
19.7695 | -104.8938
19.7732 | -104.9011
19.7769 | -104.9108
19.7805 | -104.9217
19.7842 | -104.9331
19.7879 | -104.9451
19.7916 | -104.9583
19.7953 | -104.9703
19.799 | -104.983
19.8027 | -104.9965
19.8064 | -105.0109
19.8101 | -105.0271
19.8137 | -105.0456
19.8174 | -105.0619
19.8211 | -105.0769
19.8248 | -105.0934
19.8285 | -105.11
19.8322 | -105.1274
19.8359 | -105.1476
19.8396 | -105.1671
19.8432 | -105.1855
19.8469 | -105.2027
19.8506 | -105.2156
19.8543 | -105.2278
19.858 | -105.24
19.8617 | -105.2548
19.8654 | -105.2686
19.8691 | -105.2782
19.8728 | -105.2854
19.8764 | -105.291
19.8801 | -105.2951
19.8838 | -105.299
19.8875 | -105.301
19.8912 | -105.3026
19.8949 | -105.3035
19.8986 | -105.3039
19.9023 | -105.3014
19.9059 | -105.2948
19.9096 | -105.2876
19.9133 | -105.2804
19.917 | -105.2746
19.9207 | -105.2674
19.9244 | -105.2597
19.9281 | -105.2517
19.9318 | -105.2415
19.9355 | -105.2302
19.9391 | -105.2197
19.9428 | -105.2101
19.9465 | -105.1986
19.9502 | -105.1865
19.9539 | -105.1711
19.9576 | -105.1555
19.9613 | -105.1416
19.965 | -105.1286
19.9686 | -105.1175
19.9723 | -105.1086
19.976 | -105.0986
19.9797 | -105.0889
19.9834 | -105.0788
19.9871 | -105.0696
19.9908 | -105.0621
19.9945 | -105.0543
19.9982 | -105.0475
20.0018 | -105.0417
20.0055 | -105.0324
20.0092 | -105.025
20.0129 | -105.0216
20.0166 | -105.0198
20.0203 | -105.0158
20.024 | -105.0125
20.0277 | -105.01
20.0314 | -105.0077
20.035 | -105.0054
20.0387 | -105.0032
20.0424 | -104.9992
20.0461 | -104.9964
20.0498 | -104.9964
20.0535 | -104.9971
20.0572 | -104.9975
20.0609 | -104.9966
20.0645 | -104.9969
20.0682 | -104.9975
20.0719 | -104.999
20.0756 | -105.0026
20.0793 | -105.0064
20.083 | -105.0094
20.0867 | -105.0119
20.0904 | -105.0151
20.0941 | -105.0161
20.0977 | -105.0149
20.1014 | -105.0173
20.1051 | -105.0195
20.1088 | -105.0195
20.1125 | -105.0198
20.1162 | -105.0193
20.1199 | -105.0198
20.1236 | -105.0204
20.1272 | -105.0198
20.1309 | -105.0204
20.1346 | -105.0236
20.1383 | -105.0252
20.142 | -105.0274
20.1457 | -105.0292
20.1494 | -105.0278
20.1531 | -105.0285
20.1568 | -105.0292
20.1604 | -105.0299
20.1641 | -105.0279
20.1678 | -105.0258
20.1715 | -105.0276
20.1752 | -105.0277
20.1789 | -105.0279
20.1826 | -105.0262
20.1863 | -105.0252
20.1899 | -105.0267
20.1936 | -105.0272
20.1973 | -105.0271
20.201 | -105.0247
20.2047 | -105.0228
20.2084 | -105.0239
20.2121 | -105.0225
20.2158 | -105.0182
20.2195 | -105.0115
20.2231 | -105.0067
20.2268 | -105.0025
20.2305 | -104.9984
20.2342 | -104.9957
20.2379 | -104.9935
20.2416 | -104.9919
20.2453 | -104.9908
20.249 | -104.9866
20.2527 | -104.9796
20.2563 | -104.9717
20.26 | -104.9633
20.2637 | -104.9539
20.2674 | -104.9458
20.2711 | -104.9392
20.2748 | -104.931
20.2785 | -104.9205
20.2822 | -104.9097
20.2858 | -104.8994
20.2895 | -104.8912
20.2932 | -104.8852
20.2969 | -104.8767
20.3006 | -104.8652
20.3043 | -104.8532
20.308 | -104.8442
20.3117 | -104.8355
20.3154 | -104.8267
20.319 | -104.8178
20.3227 | -104.8103
20.3264 | -104.8038
20.3301 | -104.7986
20.3338 | -104.7917
20.3375 | -104.7837
20.3412 | -104.7775
20.3449 | -104.7733
20.3485 | -104.7695
20.3522 | -104.7653
20.3559 | -104.7633
20.3596 | -104.7621
20.3633 | -104.7614
20.367 | -104.7645
20.3707 | -104.7686
20.3744 | -104.7726
20.3781 | -104.7753
20.3817 | -104.7789
20.3854 | -104.7829
20.3891 | -104.7852
20.3928 | -104.791
20.3965 | -104.7973
20.4002 | -104.8038
20.4039 | -104.8112
20.4076 | -104.8174
20.4112 | -104.8237
20.4149 | -104.8302
20.4186 | -104.837
20.4223 | -104.8437
20.426 | -104.8495
20.4297 | -104.8552
20.4334 | -104.8597
20.4371 | -104.8664
20.4408 | -104.8727
20.4444 | -104.8761
20.4481 | -104.877
20.4518 | -104.8791
20.4555 | -104.8815
20.4592 | -104.881
20.4629 | -104.8789
20.4666 | -104.8759
20.4703 | -104.8722
20.474 | -104.869
20.4776 | -104.8631
20.4813 | -104.8569
20.485 | -104.8493
20.4887 | -104.8418
20.4924 | -104.8328
20.4961 | -104.8227
20.4998 | -104.8142
20.5035 | -104.8067
20.5071 | -104.7984
20.5108 | -104.7883
20.5145 | -104.7785
20.5182 | -104.7653
20.5219 | -104.75
20.5256 | -104.7342
20.5293 | -104.7184
20.533 | -104.7043
20.5367 | -104.6909
20.5403 | -104.6791
20.544 | -104.6684
20.5477 | -104.6567
20.5514 | -104.6439
20.5551 | -104.6309
20.5588 | -104.6172
20.5625 | -104.603
20.5662 | -104.5912
20.5698 | -104.5821
20.5735 | -104.5751
20.5772 | -104.568
20.5809 | -104.5586
20.5846 | -104.5487
20.5883 | -104.541
20.592 | -104.5362
20.5957 | -104.5312
20.5994 | -104.526
20.603 | -104.5212
20.6067 | -104.518
20.6104 | -104.5159
20.6141 | -104.5141
20.6178 | -104.5147
20.6215 | -104.5161
20.6252 | -104.5181
20.6289 | -104.5196
20.6325 | -104.5199
20.6362 | -104.5207
20.6399 | -104.5242
20.6436 | -104.5285
20.6473 | -104.5322
20.651 | -104.5334
20.6547 | -104.5354
20.6584 | -104.5375
20.6621 | -104.5405
20.6657 | -104.5443
20.6694 | -104.5501
20.6731 | -104.5574
20.6768 | -104.5625
20.6805 | -104.5661
20.6842 | -104.5703
20.6879 | -104.5757
20.6916 | -104.5812
20.6953 | -104.5877
20.6989 | -104.5931
20.7026 | -104.5995
20.7063 | -104.6071
20.71 | -104.6144
20.7137 | -104.6208
20.7174 | -104.6281
20.7211 | -104.635
20.7248 | -104.6436
20.7284 | -104.6523
20.7321 | -104.6593
20.7358 | -104.6671
20.7395 | -104.6748
20.7432 | -104.6826
20.7469 | -104.6887
20.7506 | -104.6944
20.7543 | -104.7019
20.758 | -104.7111
20.7616 | -104.7202
20.7653 | -104.7283
20.769 | -104.7343
20.7727 | -104.7387
20.7764 | -104.7425
20.7801 | -104.7463
20.7838 | -104.7503
20.7875 | -104.7552
20.7911 | -104.7592
20.7948 | -104.761
20.7985 | -104.7601
20.8022 | -104.7587
20.8059 | -104.7582
20.8096 | -104.7584
20.8133 | -104.7608
20.817 | -104.7618
20.8207 | -104.7629
20.8243 | -104.7635
20.828 | -104.762
20.8317 | -104.7588
20.8354 | -104.7546
20.8391 | -104.7507
20.8428 | -104.7457
20.8465 | -104.7425
20.8502 | -104.7394
20.8539 | -104.7353
20.8575 | -104.7301
20.8612 | -104.7236
20.8649 | -104.7168
20.8686 | -104.707
20.8723 | -104.6977
20.876 | -104.69
20.8797 | -104.6809
20.8834 | -104.6695
20.887 | -104.6552
20.8907 | -104.6419
20.8944 | -104.6298
20.8981 | -104.6183
20.9018 | -104.6057
20.9055 | -104.5931
20.9092 | -104.5803
20.9129 | -104.5656
20.9166 | -104.5495
20.9202 | -104.5347
20.9239 | -104.5215
20.9276 | -104.5101
20.9313 | -104.4994
20.935 | -104.4869
20.9387 | -104.4728
20.9424 | -104.4604
20.9461 | -104.4523
20.9497 | -104.4438
20.9534 | -104.4352
20.9571 | -104.4256
20.9608 | -104.4161
20.9645 | -104.4066
20.9682 | -104.3973
20.9719 | -104.3919
20.9756 | -104.3863
20.9793 | -104.3795
20.9829 | -104.3717
20.9866 | -104.3639
20.9903 | -104.3588
20.994 | -104.3567
20.9977 | -104.3531
21.0014 | -104.3487
21.0051 | -104.3443
21.0088 | -104.3409
21.0124 | -104.3385
21.0161 | -104.3381
21.0198 | -104.3376
21.0235 | -104.3376
21.0272 | -104.3394
21.0309 | -104.3418
21.0346 | -104.3449
21.0383 | -104.3482
21.042 | -104.3511
21.0456 | -104.3528
21.0493 | -104.3545
21.053 | -104.3546
21.0567 | -104.3551
21.0604 | -104.3575
21.0641 | -104.3605
21.0678 | -104.3643
21.0715 | -104.3692
21.0752 | -104.3728
21.0788 | -104.3764
21.0825 | -104.3794
21.0862 | -104.3815
21.0899 | -104.3814
21.0936 | -104.3817
21.0973 | -104.3852
21.101 | -104.3873
21.1047 | -104.3871
21.1083 | -104.3881
21.112 | -104.3897
21.1157 | -104.3926
21.1194 | -104.395
21.1231 | -104.3976
21.1268 | -104.4006
21.1305 | -104.4029
21.1342 | -104.405
21.1379 | -104.4072
21.1415 | -104.4109
21.1452 | -104.4119
21.1489 | -104.4132
21.1526 | -104.4155
21.1563 | -104.4189
21.16 | -104.4209
21.1637 | -104.4225
21.1674 | -104.4272
21.171 | -104.4307
21.1747 | -104.4337
21.1784 | -104.4399
21.1821 | -104.4455
21.1858 | -104.4496
21.1895 | -104.453
21.1932 | -104.4561
21.1969 | -104.4575
21.2006 | -104.4596
21.2042 | -104.4622
21.2079 | -104.4648
21.2116 | -104.4697
21.2153 | -104.4769
21.219 | -104.4822
21.2227 | -104.4872
21.2264 | -104.491
21.2301 | -104.4943
21.2337 | -104.4993
21.2374 | -104.5075
21.2411 | -104.5144
21.2448 | -104.5182
21.2485 | -104.5212
21.2522 | -104.5258
21.2559 | -104.5313
21.2596 | -104.5359
21.2633 | -104.5427
21.2669 | -104.5489
21.2706 | -104.5533
21.2743 | -104.5576
21.278 | -104.5608
21.2817 | -104.5649
21.2854 | -104.5723
21.2891 | -104.5789
21.2928 | -104.5837
21.2965 | -104.5887
21.3001 | -104.5946
21.3038 | -104.6002
21.3075 | -104.6046
21.3112 | -104.6116
21.3149 | -104.6192
21.3186 | -104.6255
21.3223 | -104.6304
21.326 | -104.6355
21.3296 | -104.6398
21.3333 | -104.6464
21.337 | -104.6545
21.3407 | -104.6641
21.3444 | -104.6743
21.3481 | -104.6844
21.3518 | -104.6942
21.3555 | -104.7036
21.3592 | -104.7142
21.3628 | -104.7239
21.3665 | -104.7358
21.3702 | -104.7472
21.3739 | -104.7586
21.3776 | -104.7703
21.3813 | -104.7816
21.385 | -104.7942
21.3887 | -104.8077
21.3923 | -104.822
21.396 | -104.8369
21.3997 | -104.8495
21.4034 | -104.8621
21.4071 | -104.8749
21.4108 | -104.8871
21.4145 | -104.8966
21.4182 | -104.905
21.4219 | -104.9122
21.4255 | -104.9197
21.4292 | -104.9285
21.4329 | -104.9392
21.4366 | -104.9492
21.4403 | -104.9583
21.444 | -104.9657
21.4477 | -104.9732
21.4514 | -104.9797
21.455 | -104.9852
21.4587 | -104.9902
21.4624 | -104.9966
21.4661 | -105.0019
21.4698 | -105.005
21.4735 | -105.008
21.4772 | -105.0067
21.4809 | -105.0042
21.4846 | -105.0056
21.4882 | -105.01
21.4919 | -105.0135
21.4956 | -105.0149
21.4993 | -105.0162
21.503 | -105.0155
21.5067 | -105.0163
21.5104 | -105.0159
21.5141 | -105.0122
21.5177 | -105.0077
21.5214 | -105.004
21.5251 | -105.0002
21.5288 | -104.9971
21.5325 | -104.9946
21.5362 | -104.9934
21.5399 | -104.9914
21.5436 | -104.9872
21.5473 | -104.9847
21.5509 | -104.9855
21.5546 | -104.9841
21.5583 | -104.9815
21.562 | -104.9815
21.5657 | -104.9792
21.5694 | -104.9761
21.5731 | -104.9743
21.5768 | -104.9736
21.5805 | -104.974
21.5841 | -104.9729
21.5878 | -104.9699
21.5915 | -104.9674
21.5952 | -104.9666
21.5989 | -104.9663
21.6026 | -104.9648
21.6063 | -104.9631
21.61 | -104.9636
21.6136 | -104.9649
21.6173 | -104.962
21.621 | -104.9594
21.6247 | -104.9572
21.6284 | -104.9563
21.6321 | -104.9559
21.6358 | -104.9553
21.6395 | -104.954
21.6432 | -104.9498
21.6468 | -104.9437
21.6505 | -104.9391
21.6542 | -104.936
21.6579 | -104.9328
21.6616 | -104.9307
21.6653 | -104.9264
21.669 | -104.9215
21.6727 | -104.9143
21.6763 | -104.9085
21.68 | -104.9038
21.6837 | -104.8977
21.6874 | -104.891
21.6911 | -104.8833
21.6948 | -104.8759
21.6985 | -104.8688
21.7022 | -104.8606
21.7059 | -104.8497
21.7095 | -104.8418
21.7132 | -104.8335
21.7169 | -104.8241
21.7206 | -104.8164
21.7243 | -104.8091
21.728 | -104.8008
21.7317 | -104.7911
21.7354 | -104.7814
21.739 | -104.7713
21.7427 | -104.7593
21.7464 | -104.7493
21.7501 | -104.7389
21.7538 | -104.7276
21.7575 | -104.7159
21.7612 | -104.7039
21.7649 | -104.6902
21.7686 | -104.6762
21.7722 | -104.6646
21.7759 | -104.6539
21.7796 | -104.6422
21.7833 | -104.6328
21.787 | -104.6235
21.7907 | -104.6104
21.7944 | -104.5973
21.7981 | -104.5841
21.8018 | -104.5707
21.8054 | -104.5582
21.8091 | -104.5469
21.8128 | -104.5346
21.8165 | -104.5265
21.8202 | -104.5148
21.8239 | -104.5039
21.8276 | -104.4942
21.8313 | -104.4844
21.8349 | -104.4742
21.8386 | -104.4653
21.8423 | -104.4566
21.846 | -104.4465
21.8497 | -104.437
21.8534 | -104.4305
21.8571 | -104.423
21.8608 | -104.4157
21.8645 | -104.4097
21.8681 | -104.4055
21.8718 | -104.4001
21.8755 | -104.3957
21.8792 | -104.3932
21.8829 | -104.3896
21.8866 | -104.3884
21.8903 | -104.3879
21.894 | -104.3874
21.8976 | -104.3879
21.9013 | -104.3887
21.905 | -104.3892
21.9087 | -104.39
21.9124 | -104.391
21.9161 | -104.3938
21.9198 | -104.3961
21.9235 | -104.3982
21.9272 | -104.4026
21.9308 | -104.4076
21.9345 | -104.4121
21.9382 | -104.4157
21.9419 | -104.4208
21.9456 | -104.428
21.9493 | -104.4326
21.953 | -104.4372
21.9567 | -104.4429
21.9603 | -104.4474
21.964 | -104.4515
21.9677 | -104.4549
21.9714 | -104.4585
21.9751 | -104.4621
21.9788 | -104.4654
21.9825 | -104.47
21.9862 | -104.4738
21.9899 | -104.4763
21.9935 | -104.478
21.9972 | -104.4802
22.0009 | -104.4808
22.0046 | -104.4807
22.0083 | -104.4807
22.012 | -104.4805
22.0157 | -104.4815
22.0194 | -104.4814
22.0231 | -104.481
22.0267 | -104.4802
22.0304 | -104.4797
22.0341 | -104.4787
22.0378 | -104.4782
22.0415 | -104.4789
22.0452 | -104.4789
22.0489 | -104.4791
22.0526 | -104.4818
22.0562 | -104.484
22.0599 | -104.4853
22.0636 | -104.4869
22.0673 | -104.4887
22.071 | -104.4919
22.0747 | -104.4965
22.0784 | -104.5014
22.0821 | -104.504
22.0858 | -104.5066
22.0894 | -104.5113
22.0931 | -104.5188
22.0968 | -104.5258
22.1005 | -104.5342
22.1042 | -104.5425
22.1079 | -104.5501
22.1116 | -104.5573
22.1153 | -104.5652
22.1189 | -104.5728
22.1226 | -104.5809
22.1263 | -104.5879
22.13 | -104.5957
22.1337 | -104.6063
22.1374 | -104.6165
22.1411 | -104.6247
22.1448 | -104.6343
22.1485 | -104.6435
22.1521 | -104.6514
22.1558 | -104.6597
22.1595 | -104.6683
22.1632 | -104.6759
22.1669 | -104.682
22.1706 | -104.6881
22.1743 | -104.692
22.178 | -104.6983
22.1817 | -104.7033
22.1853 | -104.7071
22.189 | -104.7119
22.1927 | -104.7168
22.1964 | -104.7206
22.2001 | -104.7212
22.2038 | -104.7202
22.2075 | -104.7184
22.2112 | -104.7177
22.2148 | -104.7166
22.2185 | -104.7166
22.2222 | -104.7136
22.2259 | -104.7094
22.2296 | -104.7031
22.2333 | -104.6977
22.237 | -104.6912
22.2407 | -104.6825
22.2444 | -104.6738
22.248 | -104.6664
22.2517 | -104.6601
22.2554 | -104.6504
22.2591 | -104.6389
22.2628 | -104.6285
22.2665 | -104.6201
22.2702 | -104.6097
22.2739 | -104.6007
22.2775 | -104.5916
22.2812 | -104.5815
22.2849 | -104.5705
22.2886 | -104.5596
22.2923 | -104.5461
22.296 | -104.5331
22.2997 | -104.5211
22.3034 | -104.509
22.3071 | -104.4982
22.3107 | -104.4888
22.3144 | -104.4806
22.3181 | -104.4719
22.3218 | -104.4629
22.3255 | -104.4547
22.3292 | -104.4476
22.3329 | -104.4404
22.3366 | -104.4337
22.3402 | -104.4268
22.3439 | -104.4205
22.3476 | -104.4177
22.3513 | -104.4109
22.355 | -104.4048
22.3587 | -104.3998
22.3624 | -104.3944
22.3661 | -104.3888
22.3698 | -104.3858
22.3734 | -104.3824
22.3771 | -104.3811
22.3808 | -104.3822
22.3845 | -104.3827
22.3882 | -104.3811
22.3919 | -104.3793
22.3956 | -104.3786
22.3993 | -104.3797
22.403 | -104.381
22.4066 | -104.3834
22.4103 | -104.3847
22.414 | -104.3858
22.4177 | -104.3868
22.4214 | -104.3863
22.4251 | -104.388
22.4288 | -104.3915
22.4325 | -104.3947
22.4361 | -104.397
22.4398 | -104.3978
22.4435 | -104.3997
22.4472 | -104.4006
22.4509 | -104.3977
22.4546 | -104.3957
22.4583 | -104.393
22.462 | -104.3875
22.4657 | -104.3809
22.4693 | -104.3713
22.473 | -104.354
22.4767 | -104.3335
22.4804 | -104.3118
22.4841 | -104.2826
22.4878 | -104.2518
22.4915 | -104.2244
22.4952 | -104.1953
22.4988 | -104.1651
22.5025 | -104.1338
22.5062 | -104.0988
22.5099 | -104.0613
22.5136 | -104.0224
22.5173 | -103.9836
22.521 | -103.9408
22.5247 | -103.8964
22.5284 | -103.8481
22.532 | -103.7984
22.5357 | -103.7464
22.5394 | -103.6949
22.5431 | -103.6423
22.5468 | -103.5884
22.5505 | -103.532
22.5542 | -103.4745
22.5579 | -103.415
22.5615 | -103.3546
22.5652 | -103.2939
22.5689 | -103.2313
22.5726 | -103.1704
22.5763 | -103.1119
22.58 | -103.055
22.5837 | -102.9982
22.5874 | -102.9428
22.5911 | -102.887
22.5947 | -102.8338
22.5984 | -102.7842
22.6021 | -102.7373
22.6058 | -102.6935
22.6095 | -102.6542
22.6132 | -102.6172
22.6169 | -102.5838
22.6206 | -102.5556
22.6243 | -102.5307
22.6279 | -102.5085
22.6316 | -102.4894
22.6353 | -102.4748
22.639 | -102.4654
22.6427 | -102.4614
22.6464 | -102.4613
22.6501 | -102.465
22.6538 | -102.4713
22.6574 | -102.4828
22.6611 | -102.4997
22.6648 | -102.5214
22.6685 | -102.5476
22.6722 | -102.5765
22.6759 | -102.6085
22.6796 | -102.6426
22.6833 | -102.6806
22.687 | -102.7224
22.6906 | -102.7644
22.6943 | -102.8092
22.698 | -102.8557
22.7017 | -102.9033
22.7054 | -102.9534
22.7091 | -103.0048
22.7128 | -103.0569
22.7165 | -103.1087
22.7201 | -103.1606
22.7238 | -103.2124
22.7275 | -103.2636
22.7312 | -103.3139
22.7349 | -103.3625
22.7386 | -103.411
22.7423 | -103.4581
22.746 | -103.5044
22.7497 | -103.5488
22.7533 | -103.5912
22.757 | -103.6313
22.7607 | -103.6694
22.7644 | -103.7045
22.7681 | -103.7382
22.7718 | -103.7706
22.7755 | -103.8003
22.7792 | -103.8298
22.7828 | -103.8574
22.7865 | -103.8821
22.7902 | -103.9042
22.7939 | -103.9289
22.7976 | -103.9527
22.8013 | -103.9732
22.805 | -103.9958
22.8087 | -104.0147
22.8124 | -104.0255
22.816 | -104.0331
22.8197 | -104.0399
22.8234 | -104.0446
22.8271 | -104.049
22.8308 | -104.0527
22.8345 | -104.0544
22.8382 | -104.0554
22.8419 | -104.059
22.8456 | -104.0637
22.8492 | -104.0678
22.8529 | -104.069
22.8566 | -104.0701
22.8603 | -104.0709
22.864 | -104.0733
22.8677 | -104.0753
22.8714 | -104.0763
22.8751 | -104.0806
22.8787 | -104.0849
22.8824 | -104.0887
22.8861 | -104.0908
22.8898 | -104.0918
22.8935 | -104.0956
22.8972 | -104.0999
22.9009 | -104.1028
22.9046 | -104.1057
22.9083 | -104.1096
22.9119 | -104.1133
22.9156 | -104.1181
22.9193 | -104.1227
22.923 | -104.1265
22.9267 | -104.1288
22.9304 | -104.1314
22.9341 | -104.1347
22.9378 | -104.1373
22.9414 | -104.1367
22.9451 | -104.134
22.9488 | -104.1325
22.9525 | -104.1343
22.9562 | -104.1364
22.9599 | -104.137
22.9636 | -104.1353
22.9673 | -104.1328
22.971 | -104.13
22.9746 | -104.1266
22.9783 | -104.1263
22.982 | -104.1241
22.9857 | -104.1194
22.9894 | -104.1164
22.9931 | -104.113
22.9968 | -104.1102
23.0005 | -104.108
23.0041 | -104.106
23.0078 | -104.1036
23.0115 | -104.0991
23.0152 | -104.0961
23.0189 | -104.0948
23.0226 | -104.0931
23.0263 | -104.0908
23.03 | -104.0886
23.0337 | -104.0875
23.0373 | -104.0868
23.041 | -104.0864
23.0447 | -104.0849
23.0484 | -104.0848
23.0521 | -104.0838
23.0558 | -104.0826
23.0595 | -104.0823
23.0632 | -104.0826
23.0669 | -104.0826
23.0705 | -104.0827
23.0742 | -104.0835
23.0779 | -104.084
23.0816 | -104.0877
23.0853 | -104.0919
23.089 | -104.0939
23.0927 | -104.094
23.0964 | -104.0959
23.1 | -104.0972
23.1037 | -104.0966
23.1074 | -104.0964
23.1111 | -104.0984
23.1148 | -104.1016
23.1185 | -104.104
23.1222 | -104.1057
23.1259 | -104.1064
23.1296 | -104.1064
23.1332 | -104.1062
23.1369 | -104.1059
23.1406 | -104.1043
23.1443 | -104.1009
23.148 | -104.0975
23.1517 | -104.0941
23.1554 | -104.09
23.1591 | -104.0867
23.1627 | -104.0835
23.1664 | -104.0775
23.1701 | -104.0706
23.1738 | -104.0622
23.1775 | -104.0543
23.1812 | -104.0491
23.1849 | -104.0439
23.1886 | -104.0356
23.1923 | -104.0285
23.1959 | -104.0226
23.1996 | -104.0142
23.2033 | -104.0067
23.207 | -103.9989
23.2107 | -103.991
23.2144 | -103.9823
23.2181 | -103.9747
23.2218 | -103.9674
23.2254 | -103.9586
23.2291 | -103.9495
23.2328 | -103.9418
23.2365 | -103.9342
23.2402 | -103.927
23.2439 | -103.9219
23.2476 | -103.9143
23.2513 | -103.9073
23.255 | -103.9016
23.2586 | -103.8965
23.2623 | -103.8921
23.266 | -103.8896
23.2697 | -103.8888
23.2734 | -103.8888
23.2771 | -103.8892
23.2808 | -103.8885
23.2845 | -103.8879
23.2882 | -103.8882
23.2918 | -103.8904
23.2955 | -103.8914
23.2992 | -103.8954
23.3029 | -103.9017
23.3066 | -103.9066
23.3103 | -103.9134
23.314 | -103.9209
23.3177 | -103.927
23.3213 | -103.9348
23.325 | -103.945
23.3287 | -103.9539
23.3324 | -103.9632
23.3361 | -103.9729
23.3398 | -103.9838
23.3435 | -103.9951
23.3472 | -104.006
23.3509 | -104.0185
23.3545 | -104.0311
23.3582 | -104.0436
23.3619 | -104.0557
23.3656 | -104.0676
23.3693 | -104.0807
23.373 | -104.0943
23.3767 | -104.1066
23.3804 | -104.1199
23.384 | -104.1336
23.3877 | -104.149
23.3914 | -104.1628
23.3951 | -104.175
23.3988 | -104.1837
23.4025 | -104.19
23.4062 | -104.1968
23.4099 | -104.2015
23.4136 | -104.2085
23.4172 | -104.216
23.4209 | -104.2255
23.4246 | -104.2341
23.4283 | -104.2408
23.432 | -104.2474
23.4357 | -104.2542
23.4394 | -104.2593
23.4431 | -104.2613
23.4468 | -104.2636
23.4504 | -104.2671
23.4541 | -104.2694
23.4578 | -104.2703
23.4615 | -104.2702
23.4652 | -104.2672
23.4689 | -104.2667
23.4726 | -104.2661
23.4763 | -104.2659
23.4799 | -104.2671
23.4836 | -104.2665
23.4873 | -104.2666
23.491 | -104.2661
23.4947 | -104.2635
23.4984 | -104.2618
23.5021 | -104.2603
23.5058 | -104.26
23.5095 | -104.2605
23.5131 | -104.2572
23.5168 | -104.2545
23.5205 | -104.2537
23.5242 | -104.2525
23.5279 | -104.2498
23.5316 | -104.2471
23.5353 | -104.2456
23.539 | -104.2459
23.5426 | -104.2471
23.5463 | -104.2477
23.55 | -104.2466
23.5537 | -104.2438
23.5574 | -104.243
23.5611 | -104.2436
23.5648 | -104.2443
23.5685 | -104.2459
23.5722 | -104.2471
23.5758 | -104.2471
23.5795 | -104.2499
23.5832 | -104.2543
23.5869 | -104.2589
23.5906 | -104.2627
23.5943 | -104.2653
23.598 | -104.2691
23.6017 | -104.2737
23.6053 | -104.2771
23.609 | -104.2801
23.6127 | -104.2825
23.6164 | -104.2869
23.6201 | -104.2916
23.6238 | -104.2959
23.6275 | -104.3023
23.6312 | -104.3067
23.6349 | -104.3094
23.6385 | -104.314
23.6422 | -104.3174
23.6459 | -104.3202
23.6496 | -104.3243
23.6533 | -104.3277
23.657 | -104.3297
23.6607 | -104.333
23.6644 | -104.3366
23.6681 | -104.3394
23.6717 | -104.3417
23.6754 | -104.3437
23.6791 | -104.3453
23.6828 | -104.3475
23.6865 | -104.3502
23.6902 | -104.3522
23.6939 | -104.3553
23.6976 | -104.3582
23.7012 | -104.3601
23.7049 | -104.3629
23.7086 | -104.3649
23.7123 | -104.3661
23.716 | -104.3665
23.7197 | -104.3656
23.7234 | -104.3646
23.7271 | -104.3666
23.7308 | -104.3704
23.7344 | -104.3756
23.7381 | -104.3805
23.7418 | -104.3833
23.7455 | -104.3845
23.7492 | -104.388
23.7529 | -104.3913
23.7566 | -104.3933
23.7603 | -104.3951
23.7639 | -104.3974
23.7676 | -104.4016
23.7713 | -104.4056
23.775 | -104.4105
23.7787 | -104.4176
23.7824 | -104.4248
23.7861 | -104.4315
23.7898 | -104.4396
23.7935 | -104.4498
23.7971 | -104.4608
23.8008 | -104.4733
23.8045 | -104.4825
23.8082 | -104.4919
23.8119 | -104.5014
23.8156 | -104.5108
23.8193 | -104.5218
23.823 | -104.5306
23.8266 | -104.5413
23.8303 | -104.5534
23.834 | -104.5657
23.8377 | -104.578
23.8414 | -104.5908
23.8451 | -104.6021
23.8488 | -104.6144
23.8525 | -104.6259
23.8562 | -104.6372
23.8598 | -104.6487
23.8635 | -104.6611
23.8672 | -104.6727
23.8709 | -104.6815
23.8746 | -104.6895
23.8783 | -104.6957
23.882 | -104.7039
23.8857 | -104.7134
23.8894 | -104.7233
23.893 | -104.7317
23.8967 | -104.7387
23.9004 | -104.7445
23.9041 | -104.7487
23.9078 | -104.752
23.9115 | -104.756
23.9152 | -104.7597
23.9189 | -104.7599
23.9225 | -104.7594
23.9262 | -104.7595
23.9299 | -104.7596
23.9336 | -104.7573
23.9373 | -104.7552
23.941 | -104.7535
23.9447 | -104.7479
23.9484 | -104.7471
23.9521 | -104.7451
23.9557 | -104.7403
23.9594 | -104.7364
23.9631 | -104.7333
23.9668 | -104.7282
23.9705 | -104.7225
23.9742 | -104.7169
23.9779 | -104.7122
23.9816 | -104.7091
23.9852 | -104.7065
23.9889 | -104.7005
23.9926 | -104.6948
23.9963 | -104.6909
24 | -104.6872
24.0037 | -104.6818
24.0074 | -104.6768
24.0111 | -104.6735
24.0148 | -104.6699
24.0184 | -104.6658
24.0221 | -104.6615
24.0258 | -104.6578
24.0295 | -104.6545
24.0332 | -104.6513
24.0369 | -104.6487
24.0406 | -104.6454
24.0443 | -104.6418
24.0479 | -104.6397
24.0516 | -104.6388
24.0553 | -104.6383
24.059 | -104.6376
24.0627 | -104.636
24.0664 | -104.6333
24.0701 | -104.6289
24.0738 | -104.6261
24.0775 | -104.6228
24.0811 | -104.6196
24.0848 | -104.6186
24.0885 | -104.6158
24.0922 | -104.6121
24.0959 | -104.6107
24.0996 | -104.6118
24.1033 | -104.6125
24.107 | -104.6106
24.1106 | -104.6105
24.1143 | -104.6087
24.118 | -104.6045
24.1217 | -104.6018
24.1254 | -104.5989
24.1291 | -104.595
24.1328 | -104.5926
24.1365 | -104.5908
24.1402 | -104.59
24.1438 | -104.5893
24.1475 | -104.5884
24.1512 | -104.5879
24.1549 | -104.5863
24.1586 | -104.5854
24.1623 | -104.5869
24.166 | -104.5874
24.1697 | -104.5874
24.1734 | -104.5871
24.177 | -104.5872
24.1807 | -104.5863
24.1844 | -104.5864
24.1881 | -104.5868
24.1918 | -104.5878
24.1955 | -104.5888
24.1992 | -104.5902
24.2029 | -104.5912
24.2065 | -104.5912
24.2102 | -104.5934
24.2139 | -104.5959
24.2176 | -104.5989
24.2213 | -104.5994
24.225 | -104.5996
24.2287 | -104.5995
24.2324 | -104.5991
24.2361 | -104.5975
24.2397 | -104.5964
24.2434 | -104.5981
24.2471 | -104.6016
24.2508 | -104.6021
24.2545 | -104.6007
24.2582 | -104.6005
24.2619 | -104.6007
24.2656 | -104.5991
24.2692 | -104.5969
24.2729 | -104.5925
24.2766 | -104.5874
24.2803 | -104.5843
24.284 | -104.5797
24.2877 | -104.5741
24.2914 | -104.568
24.2951 | -104.5626
24.2988 | -104.5586
24.3024 | -104.5527
24.3061 | -104.5454
24.3098 | -104.5385
24.3135 | -104.53
24.3172 | -104.5227
24.3209 | -104.5142
24.3246 | -104.503
24.3283 | -104.493
24.3319 | -104.4867
24.3356 | -104.4788
24.3393 | -104.469
24.343 | -104.4596
24.3467 | -104.4513
24.3504 | -104.443
24.3541 | -104.4337
24.3578 | -104.4241
24.3615 | -104.4146
24.3651 | -104.4044
24.3688 | -104.3962
24.3725 | -104.3898
24.3762 | -104.3828
24.3799 | -104.3758
24.3836 | -104.3708
24.3873 | -104.3653
24.391 | -104.3607
24.3947 | -104.3563
24.3983 | -104.3532
24.402 | -104.3508
24.4057 | -104.3496
24.4094 | -104.3503
24.4131 | -104.3526
24.4168 | -104.3525
24.4205 | -104.3543
24.4242 | -104.3555
24.4278 | -104.3572
24.4315 | -104.3606
24.4352 | -104.3645
24.4389 | -104.3686
24.4426 | -104.371
24.4463 | -104.3747
24.45 | -104.3806
24.4537 | -104.3855
24.4574 | -104.3933
24.461 | -104.402
24.4647 | -104.4088
24.4684 | -104.4165
24.4721 | -104.4251
24.4758 | -104.4326
24.4795 | -104.4379
24.4832 | -104.4429
24.4869 | -104.4487
24.4905 | -104.456
24.4942 | -104.4639
24.4979 | -104.4708
24.5016 | -104.4758
24.5053 | -104.4813
24.509 | -104.4864
24.5127 | -104.4904
24.5164 | -104.4945
24.5201 | -104.499
24.5237 | -104.5041
24.5274 | -104.5091
24.5311 | -104.5139
24.5348 | -104.5174
24.5385 | -104.5201
24.5422 | -104.5224
24.5459 | -104.5217
24.5496 | -104.5211
24.5532 | -104.5222
24.5569 | -104.5202
24.5606 | -104.5099
24.5643 | -104.4787
24.568 | -104.4176
24.5717 | -104.32
24.5754 | -104.1588
24.5791 | -103.9871
24.5828 | -103.8553
24.5864 | -103.726
24.5901 | -103.6044
24.5938 | -103.4997
24.5975 | -103.3941
24.6012 | -103.2826
24.6049 | -103.1618
24.6086 | -103.0288
24.6123 | -102.888
24.616 | -102.7482
24.6196 | -102.606
24.6233 | -102.4628
24.627 | -102.3169
24.6307 | -102.1686
24.6344 | -102.0178
24.6381 | -101.8654
24.6418 | -101.711
24.6455 | -101.5552
24.6491 | -101.3998
24.6528 | -101.2449
24.6565 | -101.0902
24.6602 | -100.9374
24.6639 | -100.7857
24.6676 | -100.6373
24.6713 | -100.4923
24.675 | -100.3489
24.6787 | -100.2093
24.6823 | -100.0749
24.686 | -99.9448
24.6897 | -99.8194
24.6934 | -99.6993
24.6971 | -99.5845
24.7008 | -99.4758
24.7045 | -99.373
24.7082 | -99.2767
24.7118 | -99.1875
24.7155 | -99.1044
24.7192 | -99.0276
24.7229 | -98.958
24.7266 | -98.8962
24.7303 | -98.8415
24.734 | -98.7945
24.7377 | -98.7548
24.7414 | -98.7226
24.745 | -98.698
24.7487 | -98.6814
24.7524 | -98.6725
24.7561 | -98.671
24.7598 | -98.6771
24.7635 | -98.6915
24.7672 | -98.7134
24.7709 | -98.7433
24.7746 | -98.7806
24.7782 | -98.8255
24.7819 | -98.8786
24.7856 | -98.939
24.7893 | -99.0068
24.793 | -99.0814
24.7967 | -99.1628
24.8004 | -99.2506
24.8041 | -99.3446
24.8077 | -99.4453
24.8114 | -99.5522
24.8151 | -99.6646
24.8188 | -99.7821
24.8225 | -99.9058
24.8262 | -100.0341
24.8299 | -100.1679
24.8336 | -100.306
24.8373 | -100.4469
24.8409 | -100.5909
24.8446 | -100.7372
24.8483 | -100.8861
24.852 | -101.0375
24.8557 | -101.1896
24.8594 | -101.3427
24.8631 | -101.4962
24.8668 | -101.6489
24.8704 | -101.8009
24.8741 | -101.9505
24.8778 | -102.0984
24.8815 | -102.2434
24.8852 | -102.3861
24.8889 | -102.5259
24.8926 | -102.661
24.8963 | -102.7934
24.9 | -102.9213
24.9036 | -103.0439
24.9073 | -103.1621
24.911 | -103.2781
24.9147 | -103.4022
24.9184 | -103.5399
24.9221 | -103.6905
24.9258 | -103.8756
24.9295 | -104.0451
24.9331 | -104.1543
24.9368 | -104.2429
24.9405 | -104.3065
24.9442 | -104.3365
24.9479 | -104.3598
24.9516 | -104.3816
24.9553 | -104.4065
24.959 | -104.4377
24.9627 | -104.4681
24.9663 | -104.4883
24.97 | -104.5056
24.9737 | -104.5174
24.9774 | -104.5252
24.9811 | -104.5313
24.9848 | -104.5359
24.9885 | -104.5406
24.9922 | -104.5449
24.9959 | -104.5509
24.9995 | -104.5569
25.0032 | -104.5622
25.0069 | -104.5727
25.0106 | -104.5798
25.0143 | -104.5875
25.018 | -104.5943
25.0217 | -104.5999
25.0254 | -104.6065
25.029 | -104.6158
25.0327 | -104.6236
25.0364 | -104.63
25.0401 | -104.6383
25.0438 | -104.6466
25.0475 | -104.6529
25.0512 | -104.6607
25.0549 | -104.668
25.0586 | -104.6754
25.0622 | -104.6825
25.0659 | -104.6886
25.0696 | -104.6959
25.0733 | -104.7049
25.077 | -104.7137
25.0807 | -104.72
25.0844 | -104.7256
25.0881 | -104.7301
25.0917 | -104.7347
25.0954 | -104.7383
25.0991 | -104.7422
25.1028 | -104.7445
25.1065 | -104.7451
25.1102 | -104.7472
25.1139 | -104.748
25.1176 | -104.7465
25.1213 | -104.7446
25.1249 | -104.7431
25.1286 | -104.7405
25.1323 | -104.7343
25.136 | -104.7271
25.1397 | -104.7202
25.1434 | -104.7156
25.1471 | -104.7099
25.1508 | -104.7022
25.1544 | -104.6968
25.1581 | -104.6919
25.1618 | -104.6864
25.1655 | -104.6799
25.1692 | -104.6735
25.1729 | -104.6661
25.1766 | -104.659
25.1803 | -104.6505
25.184 | -104.6404
25.1876 | -104.6304
25.1913 | -104.6237
25.195 | -104.615
25.1987 | -104.607
25.2024 | -104.5993
25.2061 | -104.5925
25.2098 | -104.5847
25.2135 | -104.5757
25.2172 | -104.5655
25.2208 | -104.5558
25.2245 | -104.5475
25.2282 | -104.5388
25.2319 | -104.5312
25.2356 | -104.5236
25.2393 | -104.5145
25.243 | -104.5063
25.2467 | -104.4996
25.2503 | -104.4908
25.254 | -104.483
25.2577 | -104.4747
25.2614 | -104.4677
25.2651 | -104.4617
25.2688 | -104.4537
25.2725 | -104.4441
25.2762 | -104.4364
25.2799 | -104.4287
25.2835 | -104.4239
25.2872 | -104.4184
25.2909 | -104.4127
25.2946 | -104.4049
25.2983 | -104.3978
25.302 | -104.3917
25.3057 | -104.3866
25.3094 | -104.38
25.313 | -104.3745
25.3167 | -104.3704
25.3204 | -104.3649
25.3241 | -104.3583
25.3278 | -104.3518
25.3315 | -104.3469
25.3352 | -104.3439
25.3389 | -104.3408
25.3426 | -104.338
25.3462 | -104.3363
25.3499 | -104.3335
25.3536 | -104.3281
25.3573 | -104.3248
25.361 | -104.3204
25.3647 | -104.316
25.3684 | -104.3128
25.3721 | -104.3102
25.3757 | -104.3087
25.3794 | -104.3066
25.3831 | -104.3042
25.3868 | -104.3028
25.3905 | -104.3003
25.3942 | -104.2966
25.3979 | -104.2944
25.4016 | -104.2934
25.4053 | -104.2928
25.4089 | -104.2902
25.4126 | -104.2888
25.4163 | -104.288
25.42 | -104.2869
25.4237 | -104.2865
25.4274 | -104.2864
25.4311 | -104.2873
25.4348 | -104.2875
25.4385 | -104.2901
25.4421 | -104.2937
25.4458 | -104.298
25.4495 | -104.303
25.4532 | -104.3071
25.4569 | -104.3103
25.4606 | -104.3139
25.4643 | -104.318
25.468 | -104.321
25.4716 | -104.3243
25.4753 | -104.3281
25.479 | -104.3322
25.4827 | -104.3357
25.4864 | -104.3398
25.4901 | -104.3461
25.4938 | -104.3508
25.4975 | -104.3541
25.5012 | -104.3561
25.5048 | -104.3586
25.5085 | -104.3611
25.5122 | -104.3624
25.5159 | -104.3639
25.5196 | -104.3666
25.5233 | -104.3703
25.527 | -104.3722
25.5307 | -104.3748
25.5343 | -104.3772
25.538 | -104.3778
25.5417 | -104.3804
25.5454 | -104.383
25.5491 | -104.3849
25.5528 | -104.3857
25.5565 | -104.3885
25.5602 | -104.3905
25.5639 | -104.3929
25.5675 | -104.394
25.5712 | -104.3945
25.5749 | -104.3945
25.5786 | -104.3946
25.5823 | -104.3949
25.586 | -104.3943
25.5897 | -104.3935
25.5934 | -104.3923
25.597 | -104.3923
25.6007 | -104.3921
25.6044 | -104.3906
25.6081 | -104.3887
25.6118 | -104.3875
25.6155 | -104.3843
25.6192 | -104.3814
25.6229 | -104.3786
25.6266 | -104.3772
25.6302 | -104.3754
25.6339 | -104.3754
25.6376 | -104.3757
25.6413 | -104.3754
25.645 | -104.3743
25.6487 | -104.3717
25.6524 | -104.3677
25.6561 | -104.3653
25.6598 | -104.3643
25.6634 | -104.3626
25.6671 | -104.3607
25.6708 | -104.3603
25.6745 | -104.3593
25.6782 | -104.3595
25.6819 | -104.3605
25.6856 | -104.3619
25.6893 | -104.3619
25.6929 | -104.3614
25.6966 | -104.3616
25.7003 | -104.3615
25.704 | -104.3606
25.7077 | -104.36
25.7114 | -104.3608
25.7151 | -104.3638
25.7188 | -104.3655
25.7225 | -104.3687
25.7261 | -104.3724
25.7298 | -104.3757
25.7335 | -104.3792
25.7372 | -104.3825
25.7409 | -104.3863
25.7446 | -104.3881
25.7483 | -104.3903
25.752 | -104.3956
25.7556 | -104.4026
25.7593 | -104.4081
25.763 | -104.4126
25.7667 | -104.4186
25.7704 | -104.4257
25.7741 | -104.4325
25.7778 | -104.4388
25.7815 | -104.4455
25.7852 | -104.4535
25.7888 | -104.4589
25.7925 | -104.4648
25.7962 | -104.4722
25.7999 | -104.4769
25.8036 | -104.4822
25.8073 | -104.4863
25.811 | -104.4907
25.8147 | -104.4967
25.8183 | -104.5045
25.822 | -104.5115
25.8257 | -104.5154
25.8294 | -104.5174
25.8331 | -104.5202
25.8368 | -104.5215
25.8405 | -104.5225
25.8442 | -104.5227
25.8479 | -104.522
25.8515 | -104.5214
25.8552 | -104.5217
25.8589 | -104.5216
25.8626 | -104.5176
25.8663 | -104.5118
25.87 | -104.5086
25.8737 | -104.5064
25.8774 | -104.5007
25.8811 | -104.4919
25.8847 | -104.4826
25.8884 | -104.4743
25.8921 | -104.4638
25.8958 | -104.4526
25.8995 | -104.4421
25.9032 | -104.431
25.9069 | -104.4186
25.9106 | -104.4072
25.9142 | -104.396
25.9179 | -104.3845
25.9216 | -104.3706
25.9253 | -104.3571
25.929 | -104.3434
25.9327 | -104.3312
25.9364 | -104.3194
25.9401 | -104.3063
25.9438 | -104.294
25.9474 | -104.2818
25.9511 | -104.2675
25.9548 | -104.2521
25.9585 | -104.2354
25.9622 | -104.2198
25.9659 | -104.2055
25.9696 | -104.1919
25.9733 | -104.1796
25.9769 | -104.1679
25.9806 | -104.1568
25.9843 | -104.1474
25.988 | -104.1367
25.9917 | -104.1259
25.9954 | -104.1172
25.9991 | -104.1083
26.0028 | -104.0999
26.0065 | -104.0935
26.0101 | -104.0888
26.0138 | -104.0859
26.0175 | -104.0844
26.0212 | -104.0831
26.0249 | -104.0811
26.0286 | -104.0798
26.0323 | -104.0807
26.036 | -104.0819
26.0397 | -104.0837
26.0433 | -104.0836
26.047 | -104.0859
26.0507 | -104.0907
26.0544 | -104.0974
26.0581 | -104.1065
26.0618 | -104.1148
26.0655 | -104.125
26.0692 | -104.1353
26.0728 | -104.1487
26.0765 | -104.1591
26.0802 | -104.1694
26.0839 | -104.1809
26.0876 | -104.1924
26.0913 | -104.2024
26.095 | -104.2144
26.0987 | -104.2268
26.1024 | -104.2386
26.106 | -104.2533
26.1097 | -104.2675
26.1134 | -104.2815
26.1171 | -104.2949
26.1208 | -104.3067
26.1245 | -104.3179
26.1282 | -104.3289
26.1319 | -104.3406
26.1355 | -104.3522
26.1392 | -104.3641
26.1429 | -104.3748
26.1466 | -104.3856
26.1503 | -104.3962
26.154 | -104.4059
26.1577 | -104.4175
26.1614 | -104.4267
26.1651 | -104.4352
26.1687 | -104.4448
26.1724 | -104.4522
26.1761 | -104.4602
26.1798 | -104.4676
26.1835 | -104.4752
26.1872 | -104.4804
26.1909 | -104.4837
26.1946 | -104.4907
26.1982 | -104.4963
26.2019 | -104.4998
26.2056 | -104.5035
26.2093 | -104.5059
26.213 | -104.5081
26.2167 | -104.5111
26.2204 | -104.5131
26.2241 | -104.5131
26.2278 | -104.5126
26.2314 | -104.5148
26.2351 | -104.5169
26.2388 | -104.5152
26.2425 | -104.5108
26.2462 | -104.5082
26.2499 | -104.5081
26.2536 | -104.5073
26.2573 | -104.5054
26.261 | -104.5036
26.2646 | -104.4997
26.2683 | -104.4969
26.272 | -104.4934
26.2757 | -104.4895
26.2794 | -104.4861
26.2831 | -104.4824
26.2868 | -104.4783
26.2905 | -104.4741
26.2941 | -104.47
26.2978 | -104.4651
26.3015 | -104.4588
26.3052 | -104.4528
26.3089 | -104.4492
26.3126 | -104.4464
26.3163 | -104.4421
26.32 | -104.4376
26.3237 | -104.4336
26.3273 | -104.4299
26.331 | -104.4248
26.3347 | -104.4191
26.3384 | -104.4133
26.3421 | -104.4074
26.3458 | -104.4011
26.3495 | -104.396
26.3532 | -104.3906
26.3568 | -104.3863
26.3605 | -104.3813
26.3642 | -104.3765
26.3679 | -104.3712
26.3716 | -104.3645
26.3753 | -104.3582
26.379 | -104.3514
26.3827 | -104.347
26.3864 | -104.3433
26.39 | -104.3388
26.3937 | -104.335
26.3974 | -104.3305
26.4011 | -104.3252
26.4048 | -104.3218
26.4085 | -104.3199
26.4122 | -104.3201
26.4159 | -104.3177
26.4195 | -104.3141
26.4232 | -104.3124
26.4269 | -104.3091
26.4306 | -104.3074
26.4343 | -104.3068
26.438 | -104.307
26.4417 | -104.3082
26.4454 | -104.3094
26.4491 | -104.3099
26.4527 | -104.3091
26.4564 | -104.3071
26.4601 | -104.3058
26.4638 | -104.3072
26.4675 | -104.3068
26.4712 | -104.306
26.4749 | -104.306
26.4786 | -104.3032
26.4823 | -104.3028
26.4859 | -104.3044
26.4896 | -104.3057
26.4933 | -104.3065
26.497 | -104.3086
26.5007 | -104.3108
26.5044 | -104.3117
26.5081 | -104.3118
26.5118 | -104.3104
26.5154 | -104.3093
26.5191 | -104.3094
26.5228 | -104.3077
26.5265 | -104.3062
26.5302 | -104.305
26.5339 | -104.3045
26.5376 | -104.3038
26.5413 | -104.3043
26.545 | -104.3057
26.5486 | -104.3052
26.5523 | -104.3041
26.556 | -104.3018
26.5597 | -104.3009
26.5634 | -104.3003
26.5671 | -104.298
26.5708 | -104.2978
26.5745 | -104.2979
26.5781 | -104.2971
26.5818 | -104.2945
26.5855 | -104.2917
26.5892 | -104.2881
26.5929 | -104.2872
26.5966 | -104.2864
26.6003 | -104.2859
26.604 | -104.2859
26.6077 | -104.285
26.6113 | -104.2839
26.615 | -104.2821
26.6187 | -104.2828
26.6224 | -104.2833
26.6261 | -104.2817
26.6298 | -104.282
26.6335 | -104.2835
26.6372 | -104.2837
26.6408 | -104.2834
26.6445 | -104.2823
26.6482 | -104.2829
26.6519 | -104.2838
26.6556 | -104.2829
26.6593 | -104.2824
26.663 | -104.2825
26.6667 | -104.283
26.6704 | -104.2834
26.674 | -104.2852
26.6777 | -104.2872
26.6814 | -104.2883
26.6851 | -104.2898
26.6888 | -104.2908
26.6925 | -104.2922
26.6962 | -104.2939
26.6999 | -104.2969
26.7035 | -104.3009
26.7072 | -104.3042
26.7109 | -104.3076
26.7146 | -104.3109
26.7183 | -104.3129
26.722 | -104.3161
26.7257 | -104.3195
26.7294 | -104.322
26.7331 | -104.3241
26.7367 | -104.3278
26.7404 | -104.3308
26.7441 | -104.3339
26.7478 | -104.3394
26.7515 | -104.3437
26.7552 | -104.3486
26.7589 | -104.3522
26.7626 | -104.3558
26.7663 | -104.3601
26.7699 | -104.3646
26.7736 | -104.3707
26.7773 | -104.3774
26.781 | -104.3833
26.7847 | -104.388
26.7884 | -104.3931
26.7921 | -104.3993
26.7958 | -104.4042
26.7994 | -104.4069
26.8031 | -104.4115
26.8068 | -104.4153
26.8105 | -104.4187
26.8142 | -104.4231
26.8179 | -104.4273
26.8216 | -104.4316
26.8253 | -104.4357
26.829 | -104.441
26.8326 | -104.4472
26.8363 | -104.4531
26.84 | -104.456
26.8437 | -104.4591
26.8474 | -104.4615
26.8511 | -104.4642
26.8548 | -104.4674
26.8585 | -104.4679
26.8621 | -104.4675
26.8658 | -104.4673
26.8695 | -104.4675
26.8732 | -104.4665
26.8769 | -104.4637
26.8806 | -104.4616
26.8843 | -104.4604
26.888 | -104.4594
26.8917 | -104.4572
26.8953 | -104.4535
26.899 | -104.4506
26.9027 | -104.4465
26.9064 | -104.4419
26.9101 | -104.4377
26.9138 | -104.4341
26.9175 | -104.4288
26.9212 | -104.4234
26.9248 | -104.4199
26.9285 | -104.4157
26.9322 | -104.4113
26.9359 | -104.4065
26.9396 | -104.4032
26.9433 | -104.3984
26.947 | -104.3918
26.9507 | -104.3838
26.9544 | -104.3766
26.958 | -104.3701
26.9617 | -104.3652
26.9654 | -104.3595
26.9691 | -104.3507
26.9728 | -104.3431
26.9765 | -104.3364
26.9802 | -104.33
26.9839 | -104.3212
26.9876 | -104.3133
26.9912 | -104.3083
26.9949 | -104.3038
26.9986 | -104.2969
27.0023 | -104.291
27.006 | -104.2866
27.0097 | -104.2805
27.0134 | -104.2736
27.0171 | -104.2692
27.0207 | -104.2633
27.0244 | -104.257
27.0281 | -104.2498
27.0318 | -104.2428
27.0355 | -104.236
27.0392 | -104.2294
27.0429 | -104.2238
27.0466 | -104.2188
27.0503 | -104.2127
27.0539 | -104.2058
27.0576 | -104.1986
27.0613 | -104.1922
27.065 | -104.1859
27.0687 | -104.1787
27.0724 | -104.171
27.0761 | -104.1638
27.0798 | -104.1563
27.0834 | -104.1477
27.0871 | -104.1379
27.0908 | -104.1281
27.0945 | -104.1183
27.0982 | -104.1095
27.1019 | -104.1002
27.1056 | -104.0896
27.1093 | -104.0811
27.113 | -104.0735
27.1166 | -104.0654
27.1203 | -104.0575
27.124 | -104.0508
27.1277 | -104.0428
27.1314 | -104.0342
27.1351 | -104.0233
27.1388 | -104.0138
27.1425 | -104.0045
27.1461 | -103.9961
27.1498 | -103.9881
27.1535 | -103.9798
27.1572 | -103.9707
27.1609 | -103.9611
27.1646 | -103.9512
27.1683 | -103.943
27.172 | -103.9363
27.1757 | -103.9309
27.1793 | -103.9263
27.183 | -103.9221
27.1867 | -103.9161
27.1904 | -103.9117
27.1941 | -103.9071
27.1978 | -103.9016
27.2015 | -103.8971
27.2052 | -103.8949
27.2089 | -103.8935
27.2125 | -103.8905
27.2162 | -103.887
27.2199 | -103.8843
27.2236 | -103.8832
27.2273 | -103.8826
27.231 | -103.883
27.2347 | -103.8828
27.2384 | -103.884
27.242 | -103.885
27.2457 | -103.8872
27.2494 | -103.8888
27.2531 | -103.8891
27.2568 | -103.8908
27.2605 | -103.8939
27.2642 | -103.8974
27.2679 | -103.901
27.2716 | -103.9058
27.2752 | -103.911
27.2789 | -103.9151
27.2826 | -103.9186
27.2863 | -103.9225
27.29 | -103.9268
27.2937 | -103.9317
27.2974 | -103.9368
27.3011 | -103.9421
27.3047 | -103.9468
27.3084 | -103.951
27.3121 | -103.9557
27.3158 | -103.9609
27.3195 | -103.9648
27.3232 | -103.9669
27.3269 | -103.9719
27.3306 | -103.976
27.3343 | -103.9802
27.3379 | -103.9826
27.3416 | -103.9844
27.3453 | -103.9859
27.349 | -103.986
27.3527 | -103.9851
27.3564 | -103.9844
27.3601 | -103.9835
27.3638 | -103.9819
27.3675 | -103.9795
27.3711 | -103.9757
27.3748 | -103.9682
27.3785 | -103.9578
27.3822 | -103.9505
27.3859 | -103.942
27.3896 | -103.9301
27.3933 | -103.9169
27.397 | -103.9034
27.4006 | -103.8901
27.4043 | -103.8779
27.408 | -103.8664
27.4117 | -103.8521
27.4154 | -103.8356
27.4191 | -103.8201
27.4228 | -103.8036
27.4265 | -103.7849
27.4302 | -103.7651
27.4338 | -103.7465
27.4375 | -103.7266
27.4412 | -103.7062
27.4449 | -103.685
27.4486 | -103.6631
27.4523 | -103.6409
27.456 | -103.6187
27.4597 | -103.5967
27.4633 | -103.575
27.467 | -103.5533
27.4707 | -103.5317
27.4744 | -103.5096
27.4781 | -103.4885
27.4818 | -103.4685
27.4855 | -103.4483
27.4892 | -103.429
27.4929 | -103.4091
27.4965 | -103.3898
27.5002 | -103.3722
27.5039 | -103.3551
27.5076 | -103.3366
27.5113 | -103.3177
27.515 | -103.3008
27.5187 | -103.2853
27.5224 | -103.2722
27.526 | -103.2599
27.5297 | -103.2471
27.5334 | -103.2348
27.5371 | -103.2225
27.5408 | -103.2126
27.5445 | -103.2044
27.5482 | -103.1976
27.5519 | -103.1917
27.5556 | -103.1861
27.5592 | -103.1816
27.5629 | -103.1771
27.5666 | -103.175
27.5703 | -103.1734
27.574 | -103.1737
27.5777 | -103.1744
27.5814 | -103.1762
27.5851 | -103.1806
27.5888 | -103.1846
27.5924 | -103.1875
27.5961 | -103.1926
27.5998 | -103.1998
27.6035 | -103.2078
27.6072 | -103.216
27.6109 | -103.2254
27.6146 | -103.2359
27.6183 | -103.2472
27.6219 | -103.2592
27.6256 | -103.271
27.6293 | -103.2838
27.633 | -103.2973
27.6367 | -103.3124
27.6404 | -103.3277
27.6441 | -103.3418
27.6478 | -103.3573
27.6515 | -103.3734
27.6551 | -103.3883
27.6588 | -103.4035
27.6625 | -103.4182
27.6662 | -103.4339
27.6699 | -103.4499
27.6736 | -103.4662
27.6773 | -103.4818
27.681 | -103.4963
27.6846 | -103.5095
27.6883 | -103.5231
27.692 | -103.5361
27.6957 | -103.5486
27.6994 | -103.5627
27.7031 | -103.5758
27.7068 | -103.5881
27.7105 | -103.6013
27.7142 | -103.6137
27.7178 | -103.6237
27.7215 | -103.6338
27.7252 | -103.6442
27.7289 | -103.654
27.7326 | -103.6632
27.7363 | -103.6704
27.74 | -103.6778
27.7437 | -103.6858
27.7473 | -103.6927
27.751 | -103.6978
27.7547 | -103.7021
27.7584 | -103.7078
27.7621 | -103.7136
27.7658 | -103.7197
27.7695 | -103.7273
27.7732 | -103.7301
27.7769 | -103.7309
27.7805 | -103.733
27.7842 | -103.7354
27.7879 | -103.7369
27.7916 | -103.7387
27.7953 | -103.7396
27.799 | -103.7382
27.8027 | -103.7382
27.8064 | -103.7383
27.8101 | -103.7375
27.8137 | -103.7368
27.8174 | -103.7354
27.8211 | -103.7359
27.8248 | -103.735
27.8285 | -103.7347
27.8322 | -103.7341
27.8359 | -103.7333
27.8396 | -103.7327
27.8432 | -103.7335
27.8469 | -103.7347
27.8506 | -103.7346
27.8543 | -103.7351
27.858 | -103.7355
27.8617 | -103.7375
27.8654 | -103.7397
27.8691 | -103.7412
27.8728 | -103.7418
27.8764 | -103.7448
27.8801 | -103.7479
27.8838 | -103.7494
27.8875 | -103.7498
27.8912 | -103.7517
27.8949 | -103.756
27.8986 | -103.7627
27.9023 | -103.7698
27.9059 | -103.7769
27.9096 | -103.7843
27.9133 | -103.7889
27.917 | -103.7932
27.9207 | -103.7975
27.9244 | -103.8016
27.9281 | -103.8055
27.9318 | -103.8099
27.9355 | -103.8152
27.9391 | -103.8199
27.9428 | -103.8253
27.9465 | -103.8308
27.9502 | -103.836
27.9539 | -103.8418
27.9576 | -103.8447
27.9613 | -103.8507
27.965 | -103.8576
27.9686 | -103.863
27.9723 | -103.8684
27.976 | -103.8721
27.9797 | -103.8769
27.9834 | -103.8821
27.9871 | -103.8864
27.9908 | -103.8905
27.9945 | -103.8925
27.9982 | -103.8942
28.0018 | -103.8969