## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 241 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 255
Hood | 02-92B/P | 182
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1010.7 | hPa |  | PRC PRO 01
Humidity | 52.7 | %/Hr
Internal temperature of calibration bench | 23.82 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1014 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 23.7 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.04 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3062 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 01/06/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 06/06/2018
Technical Manager | Anthony HUE | Approval |  | 13/06/2018
CEO | Robert GLEMOT | Approval |  | 13/06/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4084 | 57.5794
0.012209 | -35.9698 | 38.4323
0.018313 | -34.9038 | 28.2492
0.024417 | -34.4381 | 21.9891
0.030521 | -34.3269 | 18.221
0.036626 | -34.1831 | 15.1601
0.04273 | -34.0871 | 12.7885
0.048834 | -34.0409 | 11.2592
0.054938 | -33.976 | 10.1123
0.061043 | -33.9152 | 9.0239
0.067147 | -33.9055 | 8.197
0.073251 | -33.907 | 7.4332
0.079355 | -33.9258 | 6.7707
0.08546 | -33.9402 | 6.1772
0.091564 | -33.9356 | 5.7373
0.097668 | -33.9178 | 5.4897
0.10377 | -33.896 | 5.2419
0.10988 | -33.879 | 4.9637
0.11598 | -33.87 | 4.6691
0.12209 | -33.8721 | 4.3834
0.12819 | -33.8825 | 4.0984
0.13429 | -33.8938 | 3.8099
0.1404 | -33.8974 | 3.5319
0.1465 | -33.8996 | 3.274
0.15261 | -33.8971 | 3.0534
0.15871 | -33.8908 | 2.8904
0.16482 | -33.8828 | 2.7865
0.17092 | -33.8807 | 2.7462
0.17702 | -33.8787 | 2.6729
0.18313 | -33.8779 | 2.5937
0.18923 | -33.8775 | 2.4839
0.19534 | -33.8763 | 2.3565
0.20144 | -33.8749 | 2.2138
0.20754 | -33.8721 | 2.0928
0.21365 | -33.8711 | 1.978
0.21975 | -33.8708 | 1.8802
0.22586 | -33.8719 | 1.7931
0.23196 | -33.8746 | 1.7213
0.23807 | -33.8788 | 1.6594
0.24417 | -33.883 | 1.6021
0.25027 | -33.8871 | 1.541
0.25638 | -33.8897 | 1.4803
0.26248 | -33.8916 | 1.4148
0.26859 | -33.8927 | 1.3507
0.27469 | -33.8931 | 1.2805
0.2808 | -33.8943 | 1.2071
0.2869 | -33.8952 | 1.1365
0.293 | -33.8962 | 1.0696
0.29911 | -33.8973 | 1.0094
0.30521 | -33.8984 | 0.95018
0.31132 | -33.8992 | 0.91125
0.31742 | -33.9006 | 0.86815
0.32353 | -33.9011 | 0.83302
0.32963 | -33.9021 | 0.80389
0.33573 | -33.9032 | 0.78361
0.34184 | -33.9047 | 0.76232
0.34794 | -33.907 | 0.73888
0.35405 | -33.9092 | 0.7151
0.36015 | -33.9118 | 0.69
0.36626 | -33.9139 | 0.66277
0.37236 | -33.9161 | 0.62632
0.37846 | -33.9184 | 0.58222
0.38457 | -33.9193 | 0.52504
0.39067 | -33.9205 | 0.47839
0.39678 | -33.9211 | 0.42639
0.40288 | -33.9216 | 0.37227
0.40899 | -33.9216 | 0.32029
0.41509 | -33.9209 | 0.26929
0.42119 | -33.9209 | 0.22595
0.4273 | -33.9209 | 0.17963
0.4334 | -33.9209 | 0.13355
0.43951 | -33.9212 | 0.090203
0.44561 | -33.9221 | 0.050538
0.45172 | -33.9237 | 0.019311
0.45782 | -33.9246 | -0.017466
0.46392 | -33.9259 | -0.053341
0.47003 | -33.9271 | -0.085247
0.47613 | -33.928 | -0.11881
0.48224 | -33.9286 | -0.15316
0.48834 | -33.9277 | -0.18737
0.49445 | -33.9276 | -0.22214
0.50055 | -33.9271 | -0.25375
0.50665 | -33.9264 | -0.28759
0.51276 | -33.9255 | -0.32131
0.51886 | -33.9245 | -0.35505
0.52497 | -33.9233 | -0.37925
0.53107 | -33.9226 | -0.4086
0.53717 | -33.922 | -0.43678
0.54328 | -33.9211 | -0.45964
0.54938 | -33.921 | -0.48218
0.55549 | -33.9211 | -0.50321
0.56159 | -33.9222 | -0.51498
0.5677 | -33.9225 | -0.52969
0.5738 | -33.9236 | -0.54609
0.5799 | -33.9245 | -0.56118
0.58601 | -33.9256 | -0.57148
0.59211 | -33.9273 | -0.5779
0.59822 | -33.9287 | -0.58756
0.60432 | -33.9303 | -0.59978
0.61043 | -33.9316 | -0.60784
0.61653 | -33.9329 | -0.61829
0.62263 | -33.9346 | -0.62779
0.62874 | -33.9353 | -0.63858
0.63484 | -33.9362 | -0.64801
0.64095 | -33.9368 | -0.65831
0.64705 | -33.9372 | -0.66816
0.65316 | -33.9375 | -0.68071
0.65926 | -33.9376 | -0.69664
0.66536 | -33.9365 | -0.71402
0.67147 | -33.9362 | -0.72724
0.67757 | -33.9359 | -0.74148
0.68368 | -33.9352 | -0.76035
0.68978 | -33.9344 | -0.77754
0.69589 | -33.9336 | -0.79635
0.70199 | -33.9326 | -0.8188
0.70809 | -33.9318 | -0.8378
0.7142 | -33.9311 | -0.85805
0.7203 | -33.93 | -0.877
0.72641 | -33.9294 | -0.89922
0.73251 | -33.9288 | -0.91891
0.73862 | -33.929 | -0.94029
0.74472 | -33.9287 | -0.95997
0.75082 | -33.9285 | -0.98172
0.75693 | -33.9288 | -1.0032
0.76303 | -33.9294 | -1.0248
0.76914 | -33.9313 | -1.0449
0.77524 | -33.9323 | -1.0641
0.78135 | -33.9336 | -1.0838
0.78745 | -33.935 | -1.1034
0.79355 | -33.9362 | -1.1229
0.79966 | -33.938 | -1.1419
0.80576 | -33.9403 | -1.1614
0.81187 | -33.9419 | -1.1827
0.81797 | -33.9439 | -1.2037
0.82408 | -33.9456 | -1.224
0.83018 | -33.9474 | -1.2444
0.83628 | -33.9493 | -1.2627
0.84239 | -33.9505 | -1.2833
0.84849 | -33.9519 | -1.3049
0.8546 | -33.9534 | -1.3247
0.8607 | -33.9547 | -1.3454
0.86681 | -33.9558 | -1.3644
0.87291 | -33.9568 | -1.3832
0.87901 | -33.9573 | -1.4013
0.88512 | -33.958 | -1.4195
0.89122 | -33.9588 | -1.4379
0.89733 | -33.9592 | -1.4536
0.90343 | -33.9597 | -1.4713
0.90953 | -33.9599 | -1.4858
0.91564 | -33.9599 | -1.5018
0.92174 | -33.9601 | -1.5168
0.92785 | -33.9603 | -1.5316
0.93395 | -33.9603 | -1.5456
0.94006 | -33.9605 | -1.5585
0.94616 | -33.9603 | -1.5704
0.95226 | -33.9604 | -1.5831
0.95837 | -33.9605 | -1.597
0.96447 | -33.9608 | -1.6099
0.97058 | -33.9606 | -1.6232
0.97668 | -33.9604 | -1.6366
0.98279 | -33.9602 | -1.6492
0.98889 | -33.96 | -1.6598
0.99499 | -33.96 | -1.6731
1.0011 | -33.9599 | -1.6837
1.0072 | -33.9595 | -1.6952
1.0133 | -33.9592 | -1.7058
1.0194 | -33.9588 | -1.7185
1.0255 | -33.9585 | -1.732
1.0316 | -33.958 | -1.7437
1.0377 | -33.9576 | -1.7572
1.0438 | -33.957 | -1.7696
1.0499 | -33.9563 | -1.7846
1.056 | -33.9556 | -1.7987
1.0621 | -33.9551 | -1.8127
1.0682 | -33.9546 | -1.8248
1.0743 | -33.954 | -1.84
1.0805 | -33.9535 | -1.8541
1.0866 | -33.9527 | -1.8689
1.0927 | -33.9518 | -1.8831
1.0988 | -33.9513 | -1.8977
1.1049 | -33.9505 | -1.9117
1.111 | -33.9497 | -1.9268
1.1171 | -33.9491 | -1.941
1.1232 | -33.9486 | -1.956
1.1293 | -33.9482 | -1.9698
1.1354 | -33.948 | -1.9846
1.1415 | -33.9477 | -1.9999
1.1476 | -33.9471 | -2.0152
1.1537 | -33.9466 | -2.0297
1.1598 | -33.9467 | -2.0444
1.1659 | -33.9465 | -2.0575
1.172 | -33.9463 | -2.0712
1.1781 | -33.9461 | -2.0851
1.1842 | -33.946 | -2.0991
1.1903 | -33.9464 | -2.1101
1.1964 | -33.9463 | -2.1234
1.2025 | -33.9464 | -2.1372
1.2086 | -33.9468 | -2.1495
1.2147 | -33.9468 | -2.1619
1.2209 | -33.9472 | -2.1729
1.227 | -33.9479 | -2.1837
1.2331 | -33.9484 | -2.1956
1.2392 | -33.9488 | -2.2074
1.2453 | -33.9492 | -2.2183
1.2514 | -33.9496 | -2.2282
1.2575 | -33.9502 | -2.2391
1.2636 | -33.9511 | -2.248
1.2697 | -33.9517 | -2.2571
1.2758 | -33.9524 | -2.2682
1.2819 | -33.9531 | -2.278
1.288 | -33.9538 | -2.2887
1.2941 | -33.9546 | -2.2971
1.3002 | -33.9553 | -2.305
1.3063 | -33.9559 | -2.3139
1.3124 | -33.9566 | -2.3228
1.3185 | -33.9573 | -2.3317
1.3246 | -33.958 | -2.3406
1.3307 | -33.9589 | -2.349
1.3368 | -33.9597 | -2.3575
1.3429 | -33.9603 | -2.3662
1.349 | -33.9608 | -2.3754
1.3551 | -33.9616 | -2.3845
1.3613 | -33.9621 | -2.3925
1.3674 | -33.9627 | -2.4022
1.3735 | -33.9632 | -2.4118
1.3796 | -33.9636 | -2.4206
1.3857 | -33.964 | -2.4296
1.3918 | -33.9645 | -2.4395
1.3979 | -33.9651 | -2.4482
1.404 | -33.9656 | -2.4592
1.4101 | -33.966 | -2.4693
1.4162 | -33.9664 | -2.4796
1.4223 | -33.9669 | -2.4907
1.4284 | -33.9674 | -2.5004
1.4345 | -33.9677 | -2.5115
1.4406 | -33.968 | -2.5246
1.4467 | -33.9684 | -2.5359
1.4528 | -33.9689 | -2.5469
1.4589 | -33.9694 | -2.5585
1.465 | -33.9696 | -2.569
1.4711 | -33.9701 | -2.58
1.4772 | -33.9706 | -2.5931
1.4833 | -33.9711 | -2.6058
1.4894 | -33.9716 | -2.6176
1.4955 | -33.9722 | -2.6293
1.5016 | -33.9727 | -2.6415
1.5078 | -33.9731 | -2.6551
1.5139 | -33.9735 | -2.6678
1.52 | -33.974 | -2.6802
1.5261 | -33.9743 | -2.6925
1.5322 | -33.9746 | -2.7062
1.5383 | -33.975 | -2.7185
1.5444 | -33.9756 | -2.7308
1.5505 | -33.9761 | -2.744
1.5566 | -33.9766 | -2.7568
1.5627 | -33.9771 | -2.7703
1.5688 | -33.9777 | -2.7837
1.5749 | -33.9781 | -2.796
1.581 | -33.9784 | -2.8091
1.5871 | -33.9788 | -2.8226
1.5932 | -33.9792 | -2.8353
1.5993 | -33.9796 | -2.8493
1.6054 | -33.98 | -2.8622
1.6115 | -33.9804 | -2.8762
1.6176 | -33.9807 | -2.8885
1.6237 | -33.9809 | -2.9018
1.6298 | -33.9812 | -2.9141
1.6359 | -33.9816 | -2.9266
1.642 | -33.9821 | -2.9383
1.6482 | -33.9823 | -2.9511
1.6543 | -33.9826 | -2.9646
1.6604 | -33.983 | -2.9764
1.6665 | -33.9831 | -2.9886
1.6726 | -33.9833 | -3.001
1.6787 | -33.9837 | -3.0133
1.6848 | -33.9836 | -3.0257
1.6909 | -33.9839 | -3.0382
1.697 | -33.9841 | -3.0501
1.7031 | -33.9842 | -3.0625
1.7092 | -33.9842 | -3.0744
1.7153 | -33.9844 | -3.0862
1.7214 | -33.9842 | -3.0982
1.7275 | -33.9841 | -3.1107
1.7336 | -33.9841 | -3.1234
1.7397 | -33.9843 | -3.1356
1.7458 | -33.9843 | -3.1476
1.7519 | -33.9843 | -3.1595
1.758 | -33.9844 | -3.1713
1.7641 | -33.9845 | -3.1827
1.7702 | -33.9845 | -3.1955
1.7763 | -33.9845 | -3.2067
1.7824 | -33.9845 | -3.219
1.7885 | -33.9845 | -3.2315
1.7947 | -33.9846 | -3.2433
1.8008 | -33.9846 | -3.2544
1.8069 | -33.9846 | -3.2657
1.813 | -33.9847 | -3.2777
1.8191 | -33.9847 | -3.2902
1.8252 | -33.9848 | -3.3028
1.8313 | -33.9849 | -3.3149
1.8374 | -33.9851 | -3.3266
1.8435 | -33.9852 | -3.3384
1.8496 | -33.9852 | -3.3513
1.8557 | -33.9854 | -3.3639
1.8618 | -33.9857 | -3.376
1.8679 | -33.9858 | -3.3879
1.874 | -33.9859 | -3.3997
1.8801 | -33.9861 | -3.412
1.8862 | -33.9863 | -3.4238
1.8923 | -33.9865 | -3.4358
1.8984 | -33.9865 | -3.4478
1.9045 | -33.9866 | -3.4592
1.9106 | -33.9867 | -3.4713
1.9167 | -33.9867 | -3.4833
1.9228 | -33.9869 | -3.4951
1.9289 | -33.987 | -3.507
1.9351 | -33.9871 | -3.5177
1.9412 | -33.9874 | -3.5288
1.9473 | -33.9877 | -3.5408
1.9534 | -33.9878 | -3.5522
1.9595 | -33.988 | -3.5639
1.9656 | -33.9883 | -3.5761
1.9717 | -33.9885 | -3.5876
1.9778 | -33.9888 | -3.599
1.9839 | -33.9891 | -3.6104
1.99 | -33.9892 | -3.6214
1.9961 | -33.9893 | -3.6336
2.0022 | -33.9895 | -3.6449
2.0083 | -33.9896 | -3.6561
2.0144 | -33.9898 | -3.6674
2.0205 | -33.9901 | -3.6791
2.0266 | -33.9904 | -3.6905
2.0327 | -33.9906 | -3.702
2.0388 | -33.991 | -3.7136
2.0449 | -33.9912 | -3.7249
2.051 | -33.9914 | -3.7362
2.0571 | -33.9918 | -3.7475
2.0632 | -33.9919 | -3.759
2.0693 | -33.9921 | -3.7711
2.0754 | -33.9924 | -3.7825
2.0816 | -33.9927 | -3.7946
2.0877 | -33.9928 | -3.8061
2.0938 | -33.9932 | -3.8182
2.0999 | -33.9934 | -3.8299
2.106 | -33.9935 | -3.8416
2.1121 | -33.9937 | -3.8535
2.1182 | -33.994 | -3.8642
2.1243 | -33.9942 | -3.875
2.1304 | -33.9944 | -3.8861
2.1365 | -33.9947 | -3.8978
2.1426 | -33.9949 | -3.9099
2.1487 | -33.9952 | -3.9211
2.1548 | -33.9956 | -3.933
2.1609 | -33.996 | -3.9445
2.167 | -33.9964 | -3.9557
2.1731 | -33.9967 | -3.9677
2.1792 | -33.9969 | -3.9788
2.1853 | -33.9973 | -3.9908
2.1914 | -33.9976 | -4.0022
2.1975 | -33.9979 | -4.0138
2.2036 | -33.9982 | -4.026
2.2097 | -33.9984 | -4.0376
2.2158 | -33.9988 | -4.0482
2.222 | -33.999 | -4.0588
2.2281 | -33.9993 | -4.0702
2.2342 | -33.9996 | -4.0815
2.2403 | -33.9998 | -4.0922
2.2464 | -34.0001 | -4.103
2.2525 | -34.0003 | -4.1144
2.2586 | -34.0007 | -4.125
2.2647 | -34.001 | -4.1355
2.2708 | -34.0013 | -4.1467
2.2769 | -34.0016 | -4.1575
2.283 | -34.0018 | -4.1686
2.2891 | -34.0021 | -4.1795
2.2952 | -34.0022 | -4.1901
2.3013 | -34.0025 | -4.2011
2.3074 | -34.0028 | -4.2116
2.3135 | -34.003 | -4.223
2.3196 | -34.0032 | -4.2333
2.3257 | -34.0034 | -4.2441
2.3318 | -34.0036 | -4.2552
2.3379 | -34.0038 | -4.266
2.344 | -34.0039 | -4.2775
2.3501 | -34.0042 | -4.2882
2.3562 | -34.0043 | -4.2986
2.3623 | -34.0045 | -4.309
2.3685 | -34.0046 | -4.3196
2.3746 | -34.0047 | -4.3309
2.3807 | -34.005 | -4.3414
2.3868 | -34.0051 | -4.3521
2.3929 | -34.0053 | -4.3628
2.399 | -34.0054 | -4.3733
2.4051 | -34.0057 | -4.3839
2.4112 | -34.0059 | -4.3945
2.4173 | -34.0061 | -4.4052
2.4234 | -34.0062 | -4.416
2.4295 | -34.0063 | -4.4269
2.4356 | -34.0065 | -4.4377
2.4417 | -34.0066 | -4.4489
2.4478 | -34.0068 | -4.4598
2.4539 | -34.007 | -4.47
2.46 | -34.0071 | -4.4815
2.4661 | -34.0073 | -4.4927
2.4722 | -34.0074 | -4.5028
2.4783 | -34.0075 | -4.5133
2.4844 | -34.0076 | -4.5244
2.4905 | -34.0077 | -4.5354
2.4966 | -34.0079 | -4.5462
2.5027 | -34.008 | -4.5567
2.5089 | -34.0081 | -4.5676
2.515 | -34.0081 | -4.5785
2.5211 | -34.0082 | -4.5892
2.5272 | -34.0083 | -4.6001
2.5333 | -34.0084 | -4.611
2.5394 | -34.0086 | -4.6218
2.5455 | -34.0087 | -4.6332
2.5516 | -34.0089 | -4.6435
2.5577 | -34.0091 | -4.6549
2.5638 | -34.0092 | -4.6666
2.5699 | -34.0093 | -4.678
2.576 | -34.0094 | -4.6895
2.5821 | -34.0095 | -4.7007
2.5882 | -34.0096 | -4.7119
2.5943 | -34.0097 | -4.7227
2.6004 | -34.0099 | -4.7336
2.6065 | -34.01 | -4.744
2.6126 | -34.0101 | -4.7544
2.6187 | -34.0103 | -4.7648
2.6248 | -34.0104 | -4.7755
2.6309 | -34.0106 | -4.7869
2.637 | -34.0108 | -4.7985
2.6431 | -34.0109 | -4.8089
2.6492 | -34.011 | -4.8202
2.6554 | -34.0112 | -4.831
2.6615 | -34.0113 | -4.8419
2.6676 | -34.0116 | -4.8533
2.6737 | -34.0118 | -4.8643
2.6798 | -34.0121 | -4.8753
2.6859 | -34.0123 | -4.8865
2.692 | -34.0124 | -4.8973
2.6981 | -34.0126 | -4.9084
2.7042 | -34.0128 | -4.9194
2.7103 | -34.013 | -4.93
2.7164 | -34.0133 | -4.9408
2.7225 | -34.0136 | -4.952
2.7286 | -34.0137 | -4.9629
2.7347 | -34.0139 | -4.9733
2.7408 | -34.0142 | -4.9832
2.7469 | -34.0144 | -4.9944
2.753 | -34.0147 | -5.0049
2.7591 | -34.015 | -5.016
2.7652 | -34.0152 | -5.0274
2.7713 | -34.0153 | -5.038
2.7774 | -34.0156 | -5.0488
2.7835 | -34.0158 | -5.0596
2.7896 | -34.016 | -5.0701
2.7958 | -34.0162 | -5.0807
2.8019 | -34.0165 | -5.091
2.808 | -34.0168 | -5.1016
2.8141 | -34.0171 | -5.1126
2.8202 | -34.0174 | -5.1233
2.8263 | -34.0176 | -5.1341
2.8324 | -34.018 | -5.1452
2.8385 | -34.0184 | -5.1559
2.8446 | -34.0187 | -5.1665
2.8507 | -34.0191 | -5.1772
2.8568 | -34.0194 | -5.1875
2.8629 | -34.0197 | -5.198
2.869 | -34.02 | -5.209
2.8751 | -34.0203 | -5.2196
2.8812 | -34.0207 | -5.2307
2.8873 | -34.021 | -5.2414
2.8934 | -34.0213 | -5.2523
2.8995 | -34.0216 | -5.2632
2.9056 | -34.0219 | -5.2743
2.9117 | -34.0222 | -5.2852
2.9178 | -34.0226 | -5.2958
2.9239 | -34.023 | -5.3073
2.93 | -34.0234 | -5.3185
2.9361 | -34.0236 | -5.3292
2.9423 | -34.024 | -5.3398
2.9484 | -34.0244 | -5.3507
2.9545 | -34.0248 | -5.3616
2.9606 | -34.0252 | -5.3729
2.9667 | -34.0255 | -5.3839
2.9728 | -34.0258 | -5.3951
2.9789 | -34.0263 | -5.4061
2.985 | -34.0267 | -5.4173
2.9911 | -34.0271 | -5.4285
2.9972 | -34.0275 | -5.4395
3.0033 | -34.0279 | -5.4509
3.0094 | -34.0283 | -5.4619
3.0155 | -34.0287 | -5.4727
3.0216 | -34.0291 | -5.4839
3.0277 | -34.0294 | -5.4955
3.0338 | -34.0298 | -5.5068
3.0399 | -34.0302 | -5.518
3.046 | -34.0305 | -5.5295
3.0521 | -34.0308 | -5.5405
3.0582 | -34.0311 | -5.5525
3.0643 | -34.0315 | -5.5639
3.0704 | -34.0318 | -5.5757
3.0765 | -34.0321 | -5.5867
3.0827 | -34.0324 | -5.5979
3.0888 | -34.0327 | -5.6091
3.0949 | -34.0331 | -5.6204
3.101 | -34.0335 | -5.6318
3.1071 | -34.0338 | -5.6435
3.1132 | -34.0341 | -5.6552
3.1193 | -34.0345 | -5.6666
3.1254 | -34.0349 | -5.6779
3.1315 | -34.0352 | -5.6898
3.1376 | -34.0355 | -5.7011
3.1437 | -34.0359 | -5.7126
3.1498 | -34.0362 | -5.7246
3.1559 | -34.0365 | -5.736
3.162 | -34.0368 | -5.7481
3.1681 | -34.037 | -5.7601
3.1742 | -34.0374 | -5.7717
3.1803 | -34.0377 | -5.7833
3.1864 | -34.0381 | -5.7945
3.1925 | -34.0384 | -5.8061
3.1986 | -34.0386 | -5.8181
3.2047 | -34.0389 | -5.8302
3.2108 | -34.0392 | -5.8413
3.2169 | -34.0395 | -5.853
3.223 | -34.0397 | -5.8647
3.2292 | -34.04 | -5.8765
3.2353 | -34.0402 | -5.8873
3.2414 | -34.0404 | -5.8987
3.2475 | -34.0406 | -5.91
3.2536 | -34.0409 | -5.9214
3.2597 | -34.0412 | -5.9324
3.2658 | -34.0414 | -5.9433
3.2719 | -34.0416 | -5.9543
3.278 | -34.0419 | -5.9658
3.2841 | -34.0421 | -5.9769
3.2902 | -34.0424 | -5.9883
3.2963 | -34.0426 | -5.9998
3.3024 | -34.0427 | -6.0109
3.3085 | -34.043 | -6.0224
3.3146 | -34.0432 | -6.0335
3.3207 | -34.0434 | -6.0451
3.3268 | -34.0436 | -6.0565
3.3329 | -34.0438 | -6.0673
3.339 | -34.044 | -6.078
3.3451 | -34.0442 | -6.0892
3.3512 | -34.0443 | -6.1
3.3573 | -34.0446 | -6.1112
3.3634 | -34.0448 | -6.1224
3.3696 | -34.045 | -6.1337
3.3757 | -34.0451 | -6.1446
3.3818 | -34.0453 | -6.1554
3.3879 | -34.0455 | -6.1665
3.394 | -34.0457 | -6.1777
3.4001 | -34.0458 | -6.1886
3.4062 | -34.0459 | -6.1995
3.4123 | -34.0461 | -6.2101
3.4184 | -34.0463 | -6.2214
3.4245 | -34.0465 | -6.2322
3.4306 | -34.0466 | -6.2429
3.4367 | -34.0468 | -6.2534
3.4428 | -34.0469 | -6.2638
3.4489 | -34.047 | -6.2748
3.455 | -34.0472 | -6.2857
3.4611 | -34.0474 | -6.2969
3.4672 | -34.0475 | -6.3079
3.4733 | -34.0477 | -6.319
3.4794 | -34.0478 | -6.3298
3.4855 | -34.048 | -6.3401
3.4916 | -34.0482 | -6.3506
3.4977 | -34.0483 | -6.3619
3.5038 | -34.0484 | -6.3727
3.5099 | -34.0486 | -6.3829
3.5161 | -34.0487 | -6.3932
3.5222 | -34.0488 | -6.4035
3.5283 | -34.0489 | -6.4143
3.5344 | -34.049 | -6.4251
3.5405 | -34.0492 | -6.4356
3.5466 | -34.0494 | -6.4465
3.5527 | -34.0495 | -6.4572
3.5588 | -34.0496 | -6.4677
3.5649 | -34.0496 | -6.4783
3.571 | -34.0497 | -6.4889
3.5771 | -34.0499 | -6.5
3.5832 | -34.05 | -6.5099
3.5893 | -34.0502 | -6.5205
3.5954 | -34.0504 | -6.5313
3.6015 | -34.0505 | -6.542
3.6076 | -34.0507 | -6.5526
3.6137 | -34.0509 | -6.5627
3.6198 | -34.0511 | -6.5729
3.6259 | -34.0513 | -6.5838
3.632 | -34.0514 | -6.5943
3.6381 | -34.0516 | -6.605
3.6442 | -34.0517 | -6.6155
3.6503 | -34.052 | -6.6256
3.6565 | -34.0522 | -6.6363
3.6626 | -34.0523 | -6.6468
3.6687 | -34.0525 | -6.6575
3.6748 | -34.0527 | -6.6678
3.6809 | -34.053 | -6.6784
3.687 | -34.0531 | -6.6893
3.6931 | -34.0534 | -6.6994
3.6992 | -34.0536 | -6.7099
3.7053 | -34.0538 | -6.7203
3.7114 | -34.0539 | -6.731
3.7175 | -34.0542 | -6.7414
3.7236 | -34.0544 | -6.7513
3.7297 | -34.0546 | -6.762
3.7358 | -34.0548 | -6.7728
3.7419 | -34.055 | -6.7836
3.748 | -34.0552 | -6.794
3.7541 | -34.0554 | -6.8045
3.7602 | -34.0557 | -6.8148
3.7663 | -34.0559 | -6.8251
3.7724 | -34.0562 | -6.8355
3.7785 | -34.0565 | -6.8458
3.7846 | -34.0567 | -6.8561
3.7907 | -34.057 | -6.8669
3.7969 | -34.0572 | -6.8775
3.803 | -34.0575 | -6.8875
3.8091 | -34.0577 | -6.898
3.8152 | -34.058 | -6.9083
3.8213 | -34.0582 | -6.919
3.8274 | -34.0585 | -6.9292
3.8335 | -34.0588 | -6.9397
3.8396 | -34.0591 | -6.9503
3.8457 | -34.0594 | -6.9608
3.8518 | -34.0598 | -6.9712
3.8579 | -34.0601 | -6.9816
3.864 | -34.0604 | -6.9922
3.8701 | -34.0607 | -7.0031
3.8762 | -34.0609 | -7.0133
3.8823 | -34.0612 | -7.0239
3.8884 | -34.0615 | -7.0344
3.8945 | -34.0619 | -7.0452
3.9006 | -34.0622 | -7.056
3.9067 | -34.0625 | -7.0669
3.9128 | -34.0628 | -7.0774
3.9189 | -34.0631 | -7.0877
3.925 | -34.0634 | -7.0983
3.9311 | -34.0636 | -7.1083
3.9372 | -34.064 | -7.1189
3.9434 | -34.0643 | -7.1295
3.9495 | -34.0647 | -7.1401
3.9556 | -34.065 | -7.1504
3.9617 | -34.0653 | -7.1605
3.9678 | -34.0657 | -7.1706
3.9739 | -34.066 | -7.1811
3.98 | -34.0662 | -7.1917
3.9861 | -34.0665 | -7.2021
3.9922 | -34.0668 | -7.2126
3.9983 | -34.0671 | -7.2227
4.0044 | -34.0675 | -7.2327
4.0105 | -34.0678 | -7.2432
4.0166 | -34.0681 | -7.2537
4.0227 | -34.0684 | -7.2642
4.0288 | -34.0687 | -7.275
4.0349 | -34.069 | -7.2852
4.041 | -34.0693 | -7.2956
4.0471 | -34.0695 | -7.3061
4.0532 | -34.0698 | -7.3162
4.0593 | -34.0702 | -7.327
4.0654 | -34.0705 | -7.3371
4.0715 | -34.0708 | -7.3481
4.0776 | -34.0712 | -7.3584
4.0838 | -34.0715 | -7.3693
4.0899 | -34.0718 | -7.3795
4.096 | -34.0721 | -7.3898
4.1021 | -34.0723 | -7.4001
4.1082 | -34.0726 | -7.4103
4.1143 | -34.0729 | -7.4204
4.1204 | -34.0733 | -7.4304
4.1265 | -34.0736 | -7.4405
4.1326 | -34.0739 | -7.4506
4.1387 | -34.0742 | -7.4608
4.1448 | -34.0745 | -7.4712
4.1509 | -34.0748 | -7.4816
4.157 | -34.075 | -7.4924
4.1631 | -34.0753 | -7.5029
4.1692 | -34.0756 | -7.5131
4.1753 | -34.0759 | -7.5235
4.1814 | -34.0762 | -7.5335
4.1875 | -34.0765 | -7.5441
4.1936 | -34.0768 | -7.5546
4.1997 | -34.0771 | -7.5646
4.2058 | -34.0773 | -7.5753
4.2119 | -34.0776 | -7.5855
4.218 | -34.0779 | -7.5956
4.2241 | -34.0781 | -7.6057
4.2303 | -34.0785 | -7.6159
4.2364 | -34.0787 | -7.6262
4.2425 | -34.0789 | -7.6365
4.2486 | -34.0792 | -7.647
4.2547 | -34.0794 | -7.6571
4.2608 | -34.0796 | -7.6675
4.2669 | -34.0798 | -7.6774
4.273 | -34.0801 | -7.6878
4.2791 | -34.0804 | -7.6982
4.2852 | -34.0806 | -7.7086
4.2913 | -34.0808 | -7.7187
4.2974 | -34.0811 | -7.7292
4.3035 | -34.0814 | -7.7397
4.3096 | -34.0816 | -7.7502
4.3157 | -34.0818 | -7.7603
4.3218 | -34.0821 | -7.7706
4.3279 | -34.0823 | -7.7809
4.334 | -34.0825 | -7.7911
4.3401 | -34.0828 | -7.801
4.3462 | -34.0831 | -7.8114
4.3523 | -34.0834 | -7.8213
4.3584 | -34.0835 | -7.8312
4.3645 | -34.0838 | -7.8413
4.3707 | -34.084 | -7.8513
4.3768 | -34.0842 | -7.8615
4.3829 | -34.0845 | -7.8721
4.389 | -34.0847 | -7.8823
4.3951 | -34.0849 | -7.8923
4.4012 | -34.0851 | -7.9028
4.4073 | -34.0854 | -7.9129
4.4134 | -34.0856 | -7.923
4.4195 | -34.0859 | -7.9332
4.4256 | -34.0862 | -7.9432
4.4317 | -34.0864 | -7.9533
4.4378 | -34.0866 | -7.9632
4.4439 | -34.0868 | -7.973
4.45 | -34.087 | -7.9832
4.4561 | -34.0871 | -7.9934
4.4622 | -34.0874 | -8.0037
4.4683 | -34.0876 | -8.0138
4.4744 | -34.0878 | -8.0242
4.4805 | -34.0881 | -8.034
4.4866 | -34.0883 | -8.0441
4.4927 | -34.0885 | -8.0545
4.4988 | -34.0887 | -8.0646
4.5049 | -34.089 | -8.0744
4.511 | -34.0892 | -8.0843
4.5172 | -34.0894 | -8.0943
4.5233 | -34.0896 | -8.1048
4.5294 | -34.0899 | -8.1147
4.5355 | -34.0901 | -8.125
4.5416 | -34.0903 | -8.1351
4.5477 | -34.0905 | -8.1452
4.5538 | -34.0907 | -8.1549
4.5599 | -34.091 | -8.165
4.566 | -34.0912 | -8.175
4.5721 | -34.0914 | -8.1851
4.5782 | -34.0917 | -8.1955
4.5843 | -34.0919 | -8.2057
4.5904 | -34.0922 | -8.2158
4.5965 | -34.0924 | -8.226
4.6026 | -34.0926 | -8.2367
4.6087 | -34.0928 | -8.2467
4.6148 | -34.093 | -8.2571
4.6209 | -34.0933 | -8.2671
4.627 | -34.0935 | -8.2776
4.6331 | -34.0937 | -8.2876
4.6392 | -34.0939 | -8.2979
4.6453 | -34.0942 | -8.308
4.6514 | -34.0944 | -8.3182
4.6576 | -34.0946 | -8.3287
4.6637 | -34.0948 | -8.3385
4.6698 | -34.095 | -8.3484
4.6759 | -34.0952 | -8.359
4.682 | -34.0954 | -8.369
4.6881 | -34.0956 | -8.3791
4.6942 | -34.0958 | -8.3892
4.7003 | -34.096 | -8.3994
4.7064 | -34.0963 | -8.4093
4.7125 | -34.0965 | -8.4198
4.7186 | -34.0967 | -8.4296
4.7247 | -34.0969 | -8.4398
4.7308 | -34.0972 | -8.4499
4.7369 | -34.0974 | -8.4601
4.743 | -34.0976 | -8.4708
4.7491 | -34.0978 | -8.4813
4.7552 | -34.098 | -8.4916
4.7613 | -34.0982 | -8.5019
4.7674 | -34.0984 | -8.5119
4.7735 | -34.0986 | -8.5225
4.7796 | -34.0989 | -8.533
4.7857 | -34.0991 | -8.5433
4.7918 | -34.0994 | -8.5533
4.7979 | -34.0996 | -8.5635
4.8041 | -34.0999 | -8.5737
4.8102 | -34.1001 | -8.5836
4.8163 | -34.1003 | -8.5937
4.8224 | -34.1005 | -8.6039
4.8285 | -34.1008 | -8.6138
4.8346 | -34.1011 | -8.6238
4.8407 | -34.1012 | -8.634
4.8468 | -34.1015 | -8.6445
4.8529 | -34.1017 | -8.6549
4.859 | -34.102 | -8.6655
4.8651 | -34.1022 | -8.6755
4.8712 | -34.1024 | -8.6855
4.8773 | -34.1026 | -8.6957
4.8834 | -34.1029 | -8.7059
4.8895 | -34.1031 | -8.7159
4.8956 | -34.1033 | -8.7261
4.9017 | -34.1035 | -8.7361
4.9078 | -34.1037 | -8.7466
4.9139 | -34.104 | -8.7565
4.92 | -34.1043 | -8.7664
4.9261 | -34.1045 | -8.7762
4.9322 | -34.1048 | -8.7864
4.9383 | -34.105 | -8.7966
4.9445 | -34.1052 | -8.8069
4.9506 | -34.1055 | -8.8169
4.9567 | -34.1057 | -8.8268
4.9628 | -34.1059 | -8.837
4.9689 | -34.1062 | -8.8475
4.975 | -34.1064 | -8.858
4.9811 | -34.1066 | -8.8679
4.9872 | -34.1069 | -8.878
4.9933 | -34.1072 | -8.8885
4.9994 | -34.1075 | -8.899
5.0055 | -34.1078 | -8.9093
5.0116 | -34.108 | -8.9199
5.0177 | -34.1083 | -8.9303
5.0238 | -34.1086 | -8.9408
5.0299 | -34.1089 | -8.9513
5.036 | -34.1092 | -8.9615
5.0421 | -34.1094 | -8.9716
5.0482 | -34.1097 | -8.9818
5.0543 | -34.1099 | -8.9918
5.0604 | -34.1103 | -9.0019
5.0665 | -34.1106 | -9.0119
5.0726 | -34.1108 | -9.0225
5.0787 | -34.1111 | -9.0329
5.0848 | -34.1114 | -9.0432
5.091 | -34.1116 | -9.0532
5.0971 | -34.1119 | -9.063
5.1032 | -34.1122 | -9.0736
5.1093 | -34.1125 | -9.0838
5.1154 | -34.1127 | -9.0938
5.1215 | -34.1129 | -9.1039
5.1276 | -34.1132 | -9.1139
5.1337 | -34.1135 | -9.1238
5.1398 | -34.1138 | -9.1337
5.1459 | -34.1141 | -9.1441
5.152 | -34.1144 | -9.1545
5.1581 | -34.1146 | -9.1649
5.1642 | -34.1149 | -9.1752
5.1703 | -34.1152 | -9.1854
5.1764 | -34.1155 | -9.1955
5.1825 | -34.1157 | -9.2054
5.1886 | -34.116 | -9.2154
5.1947 | -34.1163 | -9.2256
5.2008 | -34.1166 | -9.2359
5.2069 | -34.1168 | -9.2464
5.213 | -34.1171 | -9.2564
5.2191 | -34.1174 | -9.2667
5.2252 | -34.1177 | -9.2769
5.2314 | -34.118 | -9.2874
5.2375 | -34.1183 | -9.2973
5.2436 | -34.1185 | -9.3076
5.2497 | -34.1188 | -9.3175
5.2558 | -34.1191 | -9.3281
5.2619 | -34.1194 | -9.3385
5.268 | -34.1197 | -9.3489
5.2741 | -34.12 | -9.3592
5.2802 | -34.1203 | -9.3698
5.2863 | -34.1206 | -9.3799
5.2924 | -34.1209 | -9.3904
5.2985 | -34.1213 | -9.4007
5.3046 | -34.1216 | -9.4109
5.3107 | -34.1219 | -9.4215
5.3168 | -34.1222 | -9.4318
5.3229 | -34.1224 | -9.4421
5.329 | -34.1227 | -9.4523
5.3351 | -34.123 | -9.4623
5.3412 | -34.1233 | -9.4725
5.3473 | -34.1236 | -9.4828
5.3534 | -34.124 | -9.4933
5.3595 | -34.1242 | -9.5036
5.3656 | -34.1246 | -9.5138
5.3717 | -34.1248 | -9.5245
5.3779 | -34.1251 | -9.5351
5.384 | -34.1254 | -9.5454
5.3901 | -34.1258 | -9.5558
5.3962 | -34.126 | -9.5659
5.4023 | -34.1264 | -9.5765
5.4084 | -34.1267 | -9.5867
5.4145 | -34.127 | -9.5967
5.4206 | -34.1273 | -9.607
5.4267 | -34.1276 | -9.6176
5.4328 | -34.1279 | -9.628
5.4389 | -34.1282 | -9.638
5.445 | -34.1285 | -9.6487
5.4511 | -34.1288 | -9.6587
5.4572 | -34.1291 | -9.6691
5.4633 | -34.1294 | -9.6794
5.4694 | -34.1297 | -9.6897
5.4755 | -34.13 | -9.6998
5.4816 | -34.1303 | -9.7102
5.4877 | -34.1307 | -9.7204
5.4938 | -34.1309 | -9.7305
5.4999 | -34.1313 | -9.7407
5.506 | -34.1316 | -9.751
5.5121 | -34.1319 | -9.7613
5.5183 | -34.1322 | -9.7715
5.5244 | -34.1325 | -9.7821
5.5305 | -34.1328 | -9.7923
5.5366 | -34.1331 | -9.8024
5.5427 | -34.1334 | -9.8128
5.5488 | -34.1337 | -9.8232
5.5549 | -34.134 | -9.8337
5.561 | -34.1343 | -9.8442
5.5671 | -34.1346 | -9.8544
5.5732 | -34.1349 | -9.865
5.5793 | -34.1352 | -9.8757
5.5854 | -34.1356 | -9.886
5.5915 | -34.1359 | -9.8966
5.5976 | -34.1362 | -9.907
5.6037 | -34.1365 | -9.9174
5.6098 | -34.1367 | -9.9277
5.6159 | -34.1371 | -9.9379
5.622 | -34.1373 | -9.9484
5.6281 | -34.1377 | -9.9587
5.6342 | -34.138 | -9.9694
5.6403 | -34.1383 | -9.9798
5.6464 | -34.1386 | -9.9904
5.6525 | -34.1389 | -10.0007
5.6586 | -34.1392 | -10.0108
5.6648 | -34.1396 | -10.0212
5.6709 | -34.1399 | -10.0317
5.677 | -34.1402 | -10.0421
5.6831 | -34.1405 | -10.0525
5.6892 | -34.1408 | -10.0631
5.6953 | -34.1411 | -10.0734
5.7014 | -34.1415 | -10.0839
5.7075 | -34.1417 | -10.0943
5.7136 | -34.142 | -10.1046
5.7197 | -34.1423 | -10.115
5.7258 | -34.1426 | -10.1252
5.7319 | -34.143 | -10.1356
5.738 | -34.1433 | -10.146
5.7441 | -34.1436 | -10.1564
5.7502 | -34.144 | -10.1664
5.7563 | -34.1443 | -10.1768
5.7624 | -34.1446 | -10.1875
5.7685 | -34.1449 | -10.1979
5.7746 | -34.1452 | -10.2086
5.7807 | -34.1455 | -10.219
5.7868 | -34.1457 | -10.2296
5.7929 | -34.1461 | -10.2399
5.799 | -34.1464 | -10.2502
5.8052 | -34.1467 | -10.2604
5.8113 | -34.147 | -10.2707
5.8174 | -34.1473 | -10.281
5.8235 | -34.1476 | -10.2913
5.8296 | -34.1479 | -10.3019
5.8357 | -34.1482 | -10.3122
5.8418 | -34.1485 | -10.3228
5.8479 | -34.1487 | -10.3333
5.854 | -34.1491 | -10.3438
5.8601 | -34.1494 | -10.354
5.8662 | -34.1497 | -10.3646
5.8723 | -34.15 | -10.3748
5.8784 | -34.1503 | -10.3852
5.8845 | -34.1506 | -10.3955
5.8906 | -34.1508 | -10.4059
5.8967 | -34.1511 | -10.416
5.9028 | -34.1515 | -10.4263
5.9089 | -34.1518 | -10.4366
5.915 | -34.1521 | -10.4467
5.9211 | -34.1524 | -10.457
5.9272 | -34.1527 | -10.4672
5.9333 | -34.153 | -10.4775
5.9394 | -34.1534 | -10.4878
5.9455 | -34.1537 | -10.498
5.9517 | -34.154 | -10.5081
5.9578 | -34.1544 | -10.5187
5.9639 | -34.1547 | -10.5287
5.97 | -34.155 | -10.5391
5.9761 | -34.1554 | -10.5494
5.9822 | -34.1556 | -10.5598
5.9883 | -34.1559 | -10.5701
5.9944 | -34.1563 | -10.5806
6.0005 | -34.1566 | -10.5911
6.0066 | -34.1569 | -10.6012
6.0127 | -34.1572 | -10.6113
6.0188 | -34.1575 | -10.6219
6.0249 | -34.1578 | -10.6323
6.031 | -34.1581 | -10.6424
6.0371 | -34.1584 | -10.6528
6.0432 | -34.1587 | -10.663
6.0493 | -34.159 | -10.6731
6.0554 | -34.1594 | -10.6832
6.0615 | -34.1597 | -10.6935
6.0676 | -34.16 | -10.7041
6.0737 | -34.1604 | -10.7145
6.0798 | -34.1606 | -10.7248
6.0859 | -34.161 | -10.7351
6.0921 | -34.1613 | -10.7451
6.0982 | -34.1616 | -10.7555
6.1043 | -34.162 | -10.7656
6.1104 | -34.1623 | -10.7761
6.1165 | -34.1626 | -10.7859
6.1226 | -34.1629 | -10.7961
6.1287 | -34.1632 | -10.8064
6.1348 | -34.1636 | -10.8167
6.1409 | -34.1639 | -10.8269
6.147 | -34.1642 | -10.837
6.1531 | -34.1644 | -10.8473
6.1592 | -34.1648 | -10.8573
6.1653 | -34.1651 | -10.8674
6.1714 | -34.1654 | -10.8776
6.1775 | -34.1657 | -10.8879
6.1836 | -34.166 | -10.8983
6.1897 | -34.1663 | -10.9084
6.1958 | -34.1666 | -10.9189
6.2019 | -34.167 | -10.9292
6.208 | -34.1673 | -10.9395
6.2141 | -34.1676 | -10.9496
6.2202 | -34.1679 | -10.96
6.2263 | -34.1683 | -10.9702
6.2325 | -34.1685 | -10.9801
6.2386 | -34.1689 | -10.99
6.2447 | -34.1692 | -11.0005
6.2508 | -34.1695 | -11.0107
6.2569 | -34.1699 | -11.0209
6.263 | -34.1702 | -11.0311
6.2691 | -34.1705 | -11.0412
6.2752 | -34.1708 | -11.0515
6.2813 | -34.1711 | -11.0617
6.2874 | -34.1714 | -11.0717
6.2935 | -34.1717 | -11.0819
6.2996 | -34.1719 | -11.092
6.3057 | -34.1723 | -11.1023
6.3118 | -34.1725 | -11.1126
6.3179 | -34.1728 | -11.1227
6.324 | -34.1732 | -11.1329
6.3301 | -34.1735 | -11.143
6.3362 | -34.1738 | -11.153
6.3423 | -34.1741 | -11.1633
6.3484 | -34.1744 | -11.1735
6.3545 | -34.1747 | -11.1834
6.3606 | -34.175 | -11.1933
6.3667 | -34.1753 | -11.2033
6.3728 | -34.1756 | -11.2135
6.379 | -34.1759 | -11.2236
6.3851 | -34.1762 | -11.2335
6.3912 | -34.1765 | -11.2436
6.3973 | -34.1768 | -11.2537
6.4034 | -34.1771 | -11.2637
6.4095 | -34.1774 | -11.2738
6.4156 | -34.1777 | -11.2839
6.4217 | -34.178 | -11.294
6.4278 | -34.1783 | -11.3042
6.4339 | -34.1786 | -11.3142
6.44 | -34.1789 | -11.3242
6.4461 | -34.1792 | -11.3345
6.4522 | -34.1795 | -11.3444
6.4583 | -34.1798 | -11.3546
6.4644 | -34.1801 | -11.3648
6.4705 | -34.1804 | -11.3746
6.4766 | -34.1807 | -11.3847
6.4827 | -34.181 | -11.3949
6.4888 | -34.1813 | -11.4048
6.4949 | -34.1816 | -11.4146
6.501 | -34.1819 | -11.4248
6.5071 | -34.1822 | -11.4346
6.5132 | -34.1825 | -11.4447
6.5194 | -34.1828 | -11.4549
6.5255 | -34.1831 | -11.4647
6.5316 | -34.1834 | -11.4749
6.5377 | -34.1837 | -11.4849
6.5438 | -34.1839 | -11.4947
6.5499 | -34.1843 | -11.5047
6.556 | -34.1846 | -11.5147
6.5621 | -34.1848 | -11.525
6.5682 | -34.1852 | -11.5351
6.5743 | -34.1855 | -11.545
6.5804 | -34.1858 | -11.5551
6.5865 | -34.1861 | -11.5651
6.5926 | -34.1864 | -11.5751
6.5987 | -34.1867 | -11.5848
6.6048 | -34.187 | -11.5949
6.6109 | -34.1873 | -11.6049
6.617 | -34.1876 | -11.6151
6.6231 | -34.1879 | -11.6251
6.6292 | -34.1882 | -11.6351
6.6353 | -34.1885 | -11.6455
6.6414 | -34.1888 | -11.6554
6.6475 | -34.1891 | -11.6654
6.6536 | -34.1894 | -11.6753
6.6597 | -34.1897 | -11.685
6.6659 | -34.19 | -11.695
6.672 | -34.1903 | -11.7049
6.6781 | -34.1906 | -11.7149
6.6842 | -34.1909 | -11.7247
6.6903 | -34.1913 | -11.7347
6.6964 | -34.1916 | -11.7447
6.7025 | -34.1919 | -11.7545
6.7086 | -34.1922 | -11.7644
6.7147 | -34.1925 | -11.7745
6.7208 | -34.1928 | -11.7846
6.7269 | -34.1931 | -11.7945
6.733 | -34.1934 | -11.8042
6.7391 | -34.1937 | -11.8142
6.7452 | -34.194 | -11.8238
6.7513 | -34.1943 | -11.8335
6.7574 | -34.1946 | -11.8435
6.7635 | -34.1949 | -11.8533
6.7696 | -34.1952 | -11.8631
6.7757 | -34.1956 | -11.8731
6.7818 | -34.1959 | -11.883
6.7879 | -34.1962 | -11.8928
6.794 | -34.1965 | -11.9025
6.8001 | -34.1968 | -11.9123
6.8063 | -34.1971 | -11.9222
6.8124 | -34.1974 | -11.9322
6.8185 | -34.1977 | -11.9423
6.8246 | -34.198 | -11.952
6.8307 | -34.1983 | -11.9618
6.8368 | -34.1986 | -11.9719
6.8429 | -34.1989 | -11.9819
6.849 | -34.1992 | -11.9918
6.8551 | -34.1996 | -12.0016
6.8612 | -34.1999 | -12.0114
6.8673 | -34.2001 | -12.0213
6.8734 | -34.2005 | -12.0314
6.8795 | -34.2008 | -12.0413
6.8856 | -34.2011 | -12.0512
6.8917 | -34.2014 | -12.0611
6.8978 | -34.2017 | -12.0708
6.9039 | -34.202 | -12.0804
6.91 | -34.2024 | -12.0903
6.9161 | -34.2027 | -12.0999
6.9222 | -34.203 | -12.1098
6.9283 | -34.2033 | -12.1194
6.9344 | -34.2037 | -12.1293
6.9405 | -34.204 | -12.1392
6.9466 | -34.2043 | -12.1492
6.9528 | -34.2047 | -12.1593
6.9589 | -34.205 | -12.1693
6.965 | -34.2053 | -12.1793
6.9711 | -34.2057 | -12.1893
6.9772 | -34.206 | -12.1992
6.9833 | -34.2063 | -12.209
6.9894 | -34.2066 | -12.219
6.9955 | -34.207 | -12.2287
7.0016 | -34.2073 | -12.2385
7.0077 | -34.2076 | -12.2483
7.0138 | -34.2079 | -12.2583
7.0199 | -34.2082 | -12.2681
7.026 | -34.2086 | -12.2779
7.0321 | -34.2089 | -12.2876
7.0382 | -34.2093 | -12.2971
7.0443 | -34.2096 | -12.3072
7.0504 | -34.2099 | -12.3171
7.0565 | -34.2103 | -12.3269
7.0626 | -34.2106 | -12.3365
7.0687 | -34.2109 | -12.346
7.0748 | -34.2113 | -12.3556
7.0809 | -34.2116 | -12.3654
7.087 | -34.212 | -12.3753
7.0932 | -34.2123 | -12.3851
7.0993 | -34.2126 | -12.3949
7.1054 | -34.2129 | -12.4046
7.1115 | -34.2133 | -12.4142
7.1176 | -34.2136 | -12.424
7.1237 | -34.2139 | -12.4338
7.1298 | -34.2142 | -12.4435
7.1359 | -34.2146 | -12.4533
7.142 | -34.2149 | -12.463
7.1481 | -34.2152 | -12.4726
7.1542 | -34.2155 | -12.4825
7.1603 | -34.2159 | -12.4923
7.1664 | -34.2162 | -12.502
7.1725 | -34.2166 | -12.5119
7.1786 | -34.217 | -12.5217
7.1847 | -34.2173 | -12.5315
7.1908 | -34.2176 | -12.5415
7.1969 | -34.2179 | -12.5515
7.203 | -34.2182 | -12.5611
7.2091 | -34.2186 | -12.5708
7.2152 | -34.2189 | -12.5806
7.2213 | -34.2193 | -12.5902
7.2274 | -34.2196 | -12.5999
7.2335 | -34.2199 | -12.6096
7.2397 | -34.2203 | -12.6194
7.2458 | -34.2207 | -12.629
7.2519 | -34.221 | -12.6388
7.258 | -34.2213 | -12.6487
7.2641 | -34.2217 | -12.6585
7.2702 | -34.222 | -12.6684
7.2763 | -34.2223 | -12.6781
7.2824 | -34.2227 | -12.6881
7.2885 | -34.223 | -12.6979
7.2946 | -34.2234 | -12.7076
7.3007 | -34.2237 | -12.7175
7.3068 | -34.224 | -12.7276
7.3129 | -34.2244 | -12.7375
7.319 | -34.2248 | -12.7473
7.3251 | -34.2251 | -12.7569
7.3312 | -34.2254 | -12.7667
7.3373 | -34.2258 | -12.7762
7.3434 | -34.2261 | -12.7859
7.3495 | -34.2265 | -12.7955
7.3556 | -34.2269 | -12.8054
7.3617 | -34.2272 | -12.8153
7.3678 | -34.2276 | -12.825
7.3739 | -34.2279 | -12.835
7.3801 | -34.2282 | -12.845
7.3862 | -34.2286 | -12.8547
7.3923 | -34.229 | -12.8647
7.3984 | -34.2293 | -12.8742
7.4045 | -34.2297 | -12.8839
7.4106 | -34.2301 | -12.8937
7.4167 | -34.2304 | -12.9035
7.4228 | -34.2308 | -12.9135
7.4289 | -34.2311 | -12.9232
7.435 | -34.2314 | -12.9328
7.4411 | -34.2318 | -12.9427
7.4472 | -34.2321 | -12.9525
7.4533 | -34.2325 | -12.9621
7.4594 | -34.2329 | -12.9722
7.4655 | -34.2332 | -12.9817
7.4716 | -34.2336 | -12.9917
7.4777 | -34.2339 | -13.0014
7.4838 | -34.2343 | -13.0111
7.4899 | -34.2346 | -13.0208
7.496 | -34.2349 | -13.0304
7.5021 | -34.2353 | -13.0402
7.5082 | -34.2357 | -13.0499
7.5143 | -34.236 | -13.0597
7.5204 | -34.2364 | -13.0696
7.5266 | -34.2367 | -13.0793
7.5327 | -34.2371 | -13.089
7.5388 | -34.2374 | -13.0991
7.5449 | -34.2378 | -13.1091
7.551 | -34.2382 | -13.1186
7.5571 | -34.2385 | -13.1284
7.5632 | -34.2389 | -13.1383
7.5693 | -34.2393 | -13.148
7.5754 | -34.2396 | -13.1577
7.5815 | -34.24 | -13.1674
7.5876 | -34.2403 | -13.1775
7.5937 | -34.2407 | -13.1874
7.5998 | -34.2411 | -13.1971
7.6059 | -34.2415 | -13.2071
7.612 | -34.2418 | -13.2169
7.6181 | -34.2422 | -13.2266
7.6242 | -34.2425 | -13.2367
7.6303 | -34.2429 | -13.2467
7.6364 | -34.2432 | -13.2565
7.6425 | -34.2436 | -13.2664
7.6486 | -34.2439 | -13.2762
7.6547 | -34.2443 | -13.2859
7.6608 | -34.2447 | -13.2957
7.667 | -34.245 | -13.3057
7.6731 | -34.2454 | -13.3156
7.6792 | -34.2457 | -13.3256
7.6853 | -34.2461 | -13.3355
7.6914 | -34.2465 | -13.3454
7.6975 | -34.2468 | -13.3553
7.7036 | -34.2472 | -13.3652
7.7097 | -34.2476 | -13.3752
7.7158 | -34.2479 | -13.3851
7.7219 | -34.2483 | -13.3947
7.728 | -34.2487 | -13.4046
7.7341 | -34.249 | -13.4146
7.7402 | -34.2494 | -13.4245
7.7463 | -34.2497 | -13.4344
7.7524 | -34.2501 | -13.4445
7.7585 | -34.2505 | -13.4544
7.7646 | -34.2508 | -13.4642
7.7707 | -34.2512 | -13.4742
7.7768 | -34.2515 | -13.4842
7.7829 | -34.2519 | -13.4943
7.789 | -34.2522 | -13.5044
7.7951 | -34.2526 | -13.5141
7.8012 | -34.253 | -13.524
7.8073 | -34.2534 | -13.5341
7.8135 | -34.2537 | -13.5439
7.8196 | -34.2541 | -13.554
7.8257 | -34.2545 | -13.5641
7.8318 | -34.2548 | -13.574
7.8379 | -34.2552 | -13.5838
7.844 | -34.2555 | -13.5937
7.8501 | -34.2559 | -13.6037
7.8562 | -34.2563 | -13.6138
7.8623 | -34.2567 | -13.6235
7.8684 | -34.257 | -13.6335
7.8745 | -34.2574 | -13.6435
7.8806 | -34.2577 | -13.6536
7.8867 | -34.2581 | -13.6632
7.8928 | -34.2585 | -13.6729
7.8989 | -34.2588 | -13.6828
7.905 | -34.2592 | -13.6928
7.9111 | -34.2595 | -13.7028
7.9172 | -34.2599 | -13.7126
7.9233 | -34.2603 | -13.7226
7.9294 | -34.2606 | -13.7326
7.9355 | -34.261 | -13.7423
7.9416 | -34.2613 | -13.7522
7.9477 | -34.2617 | -13.7619
7.9539 | -34.2621 | -13.7718
7.96 | -34.2624 | -13.7819
7.9661 | -34.2628 | -13.7918
7.9722 | -34.2632 | -13.8016
7.9783 | -34.2636 | -13.8118
7.9844 | -34.264 | -13.8216
7.9905 | -34.2643 | -13.8317
7.9966 | -34.2647 | -13.8418
8.0027 | -34.265 | -13.8518
8.0088 | -34.2654 | -13.8617
8.0149 | -34.2657 | -13.8717
8.021 | -34.2661 | -13.8816
8.0271 | -34.2665 | -13.8919
8.0332 | -34.2668 | -13.9017
8.0393 | -34.2672 | -13.9117
8.0454 | -34.2676 | -13.9219
8.0515 | -34.268 | -13.9319
8.0576 | -34.2683 | -13.9421
8.0637 | -34.2687 | -13.952
8.0698 | -34.269 | -13.9617
8.0759 | -34.2694 | -13.9715
8.082 | -34.2697 | -13.9814
8.0881 | -34.2701 | -13.9914
8.0942 | -34.2705 | -14.0014
8.1004 | -34.2709 | -14.0114
8.1065 | -34.2712 | -14.0211
8.1126 | -34.2715 | -14.0311
8.1187 | -34.2719 | -14.0411
8.1248 | -34.2723 | -14.0508
8.1309 | -34.2727 | -14.0605
8.137 | -34.273 | -14.0702
8.1431 | -34.2734 | -14.0803
8.1492 | -34.2737 | -14.0902
8.1553 | -34.2741 | -14.1002
8.1614 | -34.2745 | -14.1102
8.1675 | -34.2748 | -14.1201
8.1736 | -34.2751 | -14.1302
8.1797 | -34.2755 | -14.1402
8.1858 | -34.2759 | -14.1502
8.1919 | -34.2762 | -14.1601
8.198 | -34.2766 | -14.1698
8.2041 | -34.277 | -14.1798
8.2102 | -34.2773 | -14.1897
8.2163 | -34.2777 | -14.1994
8.2224 | -34.2781 | -14.2092
8.2285 | -34.2784 | -14.2189
8.2346 | -34.2788 | -14.2288
8.2408 | -34.2791 | -14.2385
8.2469 | -34.2795 | -14.2485
8.253 | -34.2799 | -14.2583
8.2591 | -34.2803 | -14.2682
8.2652 | -34.2807 | -14.2781
8.2713 | -34.281 | -14.288
8.2774 | -34.2814 | -14.298
8.2835 | -34.2818 | -14.3082
8.2896 | -34.2821 | -14.3179
8.2957 | -34.2825 | -14.3278
8.3018 | -34.2829 | -14.3377
8.3079 | -34.2832 | -14.3475
8.314 | -34.2836 | -14.3573
8.3201 | -34.284 | -14.3672
8.3262 | -34.2843 | -14.3772
8.3323 | -34.2847 | -14.387
8.3384 | -34.285 | -14.397
8.3445 | -34.2854 | -14.4068
8.3506 | -34.2858 | -14.4167
8.3567 | -34.2862 | -14.4267
8.3628 | -34.2865 | -14.4366
8.3689 | -34.2869 | -14.4463
8.375 | -34.2873 | -14.4563
8.3812 | -34.2876 | -14.4661
8.3873 | -34.288 | -14.476
8.3934 | -34.2883 | -14.486
8.3995 | -34.2887 | -14.4957
8.4056 | -34.289 | -14.5056
8.4117 | -34.2894 | -14.5155
8.4178 | -34.2897 | -14.5254
8.4239 | -34.2901 | -14.5352
8.43 | -34.2905 | -14.5451
8.4361 | -34.2908 | -14.555
8.4422 | -34.2912 | -14.5647
8.4483 | -34.2915 | -14.5745
8.4544 | -34.2919 | -14.5844
8.4605 | -34.2923 | -14.5942
8.4666 | -34.2926 | -14.6043
8.4727 | -34.293 | -14.6142
8.4788 | -34.2933 | -14.624
8.4849 | -34.2937 | -14.6338
8.491 | -34.2941 | -14.6436
8.4971 | -34.2944 | -14.6532
8.5032 | -34.2948 | -14.6629
8.5093 | -34.2952 | -14.6726
8.5154 | -34.2955 | -14.6825
8.5215 | -34.2959 | -14.6923
8.5277 | -34.2962 | -14.7021
8.5338 | -34.2966 | -14.7118
8.5399 | -34.297 | -14.7216
8.546 | -34.2974 | -14.7312
8.5521 | -34.2977 | -14.741
8.5582 | -34.2981 | -14.7508
8.5643 | -34.2984 | -14.7604
8.5704 | -34.2988 | -14.7699
8.5765 | -34.2992 | -14.7799
8.5826 | -34.2995 | -14.7894
8.5887 | -34.2999 | -14.7991
8.5948 | -34.3003 | -14.8089
8.6009 | -34.3007 | -14.8185
8.607 | -34.301 | -14.8284
8.6131 | -34.3014 | -14.8383
8.6192 | -34.3018 | -14.8478
8.6253 | -34.3022 | -14.8575
8.6314 | -34.3025 | -14.8672
8.6375 | -34.3028 | -14.877
8.6436 | -34.3032 | -14.8868
8.6497 | -34.3036 | -14.8964
8.6558 | -34.3039 | -14.9061
8.6619 | -34.3042 | -14.9159
8.6681 | -34.3046 | -14.9257
8.6742 | -34.305 | -14.9354
8.6803 | -34.3053 | -14.9452
8.6864 | -34.3057 | -14.9546
8.6925 | -34.3061 | -14.9642
8.6986 | -34.3064 | -14.974
8.7047 | -34.3068 | -14.984
8.7108 | -34.3071 | -14.9938
8.7169 | -34.3075 | -15.0035
8.723 | -34.3078 | -15.013
8.7291 | -34.3082 | -15.0226
8.7352 | -34.3086 | -15.0325
8.7413 | -34.3089 | -15.0423
8.7474 | -34.3092 | -15.0518
8.7535 | -34.3096 | -15.0616
8.7596 | -34.31 | -15.0713
8.7657 | -34.3104 | -15.0812
8.7718 | -34.3107 | -15.091
8.7779 | -34.3111 | -15.1008
8.784 | -34.3115 | -15.1105
8.7901 | -34.3119 | -15.1203
8.7962 | -34.3122 | -15.1297
8.8023 | -34.3126 | -15.1394
8.8084 | -34.3129 | -15.149
8.8146 | -34.3133 | -15.1587
8.8207 | -34.3136 | -15.1684
8.8268 | -34.314 | -15.1777
8.8329 | -34.3144 | -15.1874
8.839 | -34.3148 | -15.1971
8.8451 | -34.3151 | -15.2067
8.8512 | -34.3155 | -15.2162
8.8573 | -34.3158 | -15.226
8.8634 | -34.3162 | -15.2354
8.8695 | -34.3166 | -15.2452
8.8756 | -34.3169 | -15.2549
8.8817 | -34.3173 | -15.2644
8.8878 | -34.3177 | -15.2743
8.8939 | -34.318 | -15.284
8.9 | -34.3184 | -15.2936
8.9061 | -34.3187 | -15.3033
8.9122 | -34.3191 | -15.3129
8.9183 | -34.3195 | -15.3226
8.9244 | -34.3199 | -15.3324
8.9305 | -34.3203 | -15.3418
8.9366 | -34.3206 | -15.3514
8.9427 | -34.321 | -15.3611
8.9488 | -34.3214 | -15.3707
8.955 | -34.3217 | -15.3803
8.9611 | -34.3221 | -15.3901
8.9672 | -34.3225 | -15.3997
8.9733 | -34.3229 | -15.4091
8.9794 | -34.3233 | -15.4186
8.9855 | -34.3237 | -15.4284
8.9916 | -34.324 | -15.4382
8.9977 | -34.3244 | -15.4478
9.0038 | -34.3248 | -15.4573
9.0099 | -34.3252 | -15.4671
9.016 | -34.3256 | -15.4768
9.0221 | -34.3259 | -15.4862
9.0282 | -34.3263 | -15.4957
9.0343 | -34.3267 | -15.5053
9.0404 | -34.3271 | -15.5149
9.0465 | -34.3275 | -15.5247
9.0526 | -34.3279 | -15.5342
9.0587 | -34.3283 | -15.5439
9.0648 | -34.3287 | -15.5534
9.0709 | -34.329 | -15.563
9.077 | -34.3295 | -15.5726
9.0831 | -34.3298 | -15.5825
9.0892 | -34.3302 | -15.5922
9.0953 | -34.3306 | -15.6018
9.1015 | -34.331 | -15.6114
9.1076 | -34.3314 | -15.6208
9.1137 | -34.3317 | -15.6303
9.1198 | -34.3321 | -15.6399
9.1259 | -34.3325 | -15.6495
9.132 | -34.3329 | -15.6592
9.1381 | -34.3333 | -15.6687
9.1442 | -34.3337 | -15.678
9.1503 | -34.3341 | -15.6877
9.1564 | -34.3345 | -15.6973
9.1625 | -34.3348 | -15.7069
9.1686 | -34.3352 | -15.7164
9.1747 | -34.3356 | -15.7261
9.1808 | -34.336 | -15.7357
9.1869 | -34.3364 | -15.7452
9.193 | -34.3368 | -15.7549
9.1991 | -34.3372 | -15.7645
9.2052 | -34.3376 | -15.774
9.2113 | -34.338 | -15.7836
9.2174 | -34.3384 | -15.7933
9.2235 | -34.3388 | -15.8028
9.2296 | -34.3392 | -15.8122
9.2357 | -34.3396 | -15.8219
9.2419 | -34.34 | -15.8314
9.248 | -34.3404 | -15.841
9.2541 | -34.3408 | -15.8506
9.2602 | -34.3412 | -15.8605
9.2663 | -34.3416 | -15.8697
9.2724 | -34.342 | -15.8791
9.2785 | -34.3424 | -15.8887
9.2846 | -34.3428 | -15.898
9.2907 | -34.3432 | -15.9076
9.2968 | -34.3436 | -15.9172
9.3029 | -34.344 | -15.9268
9.309 | -34.3444 | -15.9365
9.3151 | -34.3448 | -15.9459
9.3212 | -34.3452 | -15.9554
9.3273 | -34.3456 | -15.9649
9.3334 | -34.346 | -15.9743
9.3395 | -34.3464 | -15.9837
9.3456 | -34.3467 | -15.9931
9.3517 | -34.3471 | -16.0027
9.3578 | -34.3475 | -16.0122
9.3639 | -34.3479 | -16.0217
9.37 | -34.3483 | -16.0313
9.3761 | -34.3487 | -16.0409
9.3822 | -34.3491 | -16.0504
9.3884 | -34.3496 | -16.0596
9.3945 | -34.3499 | -16.0693
9.4006 | -34.3504 | -16.0788
9.4067 | -34.3508 | -16.0884
9.4128 | -34.3512 | -16.098
9.4189 | -34.3516 | -16.1078
9.425 | -34.352 | -16.1171
9.4311 | -34.3524 | -16.1266
9.4372 | -34.3528 | -16.1361
9.4433 | -34.3532 | -16.1456
9.4494 | -34.3536 | -16.1549
9.4555 | -34.3541 | -16.1645
9.4616 | -34.3545 | -16.174
9.4677 | -34.3549 | -16.1835
9.4738 | -34.3553 | -16.1931
9.4799 | -34.3557 | -16.2026
9.486 | -34.3561 | -16.2122
9.4921 | -34.3565 | -16.2217
9.4982 | -34.3569 | -16.2309
9.5043 | -34.3573 | -16.2404
9.5104 | -34.3577 | -16.2497
9.5165 | -34.3581 | -16.2592
9.5226 | -34.3585 | -16.2689
9.5288 | -34.3589 | -16.2782
9.5349 | -34.3593 | -16.2878
9.541 | -34.3597 | -16.2971
9.5471 | -34.3601 | -16.3066
9.5532 | -34.3605 | -16.3162
9.5593 | -34.3609 | -16.3259
9.5654 | -34.3613 | -16.3353
9.5715 | -34.3617 | -16.3447
9.5776 | -34.3621 | -16.3544
9.5837 | -34.3625 | -16.364
9.5898 | -34.3629 | -16.3736
9.5959 | -34.3633 | -16.3831
9.602 | -34.3637 | -16.3925
9.6081 | -34.3641 | -16.4019
9.6142 | -34.3645 | -16.4112
9.6203 | -34.3649 | -16.4207
9.6264 | -34.3653 | -16.4299
9.6325 | -34.3657 | -16.4395
9.6386 | -34.3661 | -16.4488
9.6447 | -34.3665 | -16.4583
9.6508 | -34.3669 | -16.4675
9.6569 | -34.3673 | -16.477
9.663 | -34.3677 | -16.4866
9.6691 | -34.3681 | -16.4959
9.6753 | -34.3685 | -16.5053
9.6814 | -34.3689 | -16.5147
9.6875 | -34.3693 | -16.5239
9.6936 | -34.3697 | -16.5332
9.6997 | -34.3702 | -16.5429
9.7058 | -34.3706 | -16.5523
9.7119 | -34.371 | -16.5617
9.718 | -34.3714 | -16.5712
9.7241 | -34.3718 | -16.5805
9.7302 | -34.3722 | -16.59
9.7363 | -34.3726 | -16.5994
9.7424 | -34.3731 | -16.6088
9.7485 | -34.3735 | -16.6182
9.7546 | -34.3739 | -16.6275
9.7607 | -34.3743 | -16.6372
9.7668 | -34.3747 | -16.6468
9.7729 | -34.3751 | -16.6562
9.779 | -34.3756 | -16.6654
9.7851 | -34.376 | -16.6749
9.7912 | -34.3765 | -16.6842
9.7973 | -34.3769 | -16.6939
9.8034 | -34.3773 | -16.703
9.8095 | -34.3777 | -16.7127
9.8157 | -34.3781 | -16.7222
9.8218 | -34.3786 | -16.7318
9.8279 | -34.379 | -16.7411
9.834 | -34.3793 | -16.7507
9.8401 | -34.3797 | -16.7601
9.8462 | -34.3802 | -16.7692
9.8523 | -34.3806 | -16.7787
9.8584 | -34.381 | -16.7882
9.8645 | -34.3814 | -16.7977
9.8706 | -34.3818 | -16.8072
9.8767 | -34.3821 | -16.8166
9.8828 | -34.3826 | -16.8258
9.8889 | -34.383 | -16.835
9.895 | -34.3834 | -16.8444
9.9011 | -34.3838 | -16.8538
9.9072 | -34.3842 | -16.8632
9.9133 | -34.3846 | -16.8726
9.9194 | -34.385 | -16.8818
9.9255 | -34.3854 | -16.8911
9.9316 | -34.3858 | -16.9005
9.9377 | -34.3863 | -16.91
9.9438 | -34.3867 | -16.9193
9.9499 | -34.3871 | -16.929
9.956 | -34.3875 | -16.9384
9.9622 | -34.3879 | -16.9478
9.9683 | -34.3883 | -16.9572
9.9744 | -34.3887 | -16.9667
9.9805 | -34.3892 | -16.9761
9.9866 | -34.3896 | -16.9855
9.9927 | -34.39 | -16.9946
9.9988 | -34.3904 | -17.0041
10.0049 | -34.3908 | -17.0135
10.011 | -34.3912 | -17.0229
10.0171 | -34.3916 | -17.0321
10.0232 | -34.3921 | -17.0414
10.0293 | -34.3925 | -17.0509
10.0354 | -34.3929 | -17.0603
10.0415 | -34.3934 | -17.0697
10.0476 | -34.3938 | -17.0792
10.0537 | -34.3942 | -17.0887
10.0598 | -34.3946 | -17.0981
10.0659 | -34.395 | -17.1074
10.072 | -34.3954 | -17.1166
10.0781 | -34.3958 | -17.126
10.0842 | -34.3962 | -17.1353
10.0903 | -34.3966 | -17.1446
10.0964 | -34.397 | -17.1538
10.1026 | -34.3974 | -17.1631
10.1087 | -34.3978 | -17.1724
10.1148 | -34.3982 | -17.1819
10.1209 | -34.3986 | -17.1911
10.127 | -34.399 | -17.2003
10.1331 | -34.3995 | -17.2095
10.1392 | -34.3999 | -17.2186
10.1453 | -34.4003 | -17.2279
10.1514 | -34.4007 | -17.2374
10.1575 | -34.4011 | -17.2469
10.1636 | -34.4016 | -17.2558
10.1697 | -34.402 | -17.2652
10.1758 | -34.4024 | -17.2743
10.1819 | -34.4028 | -17.2836
10.188 | -34.4032 | -17.293
10.1941 | -34.4036 | -17.3023
10.2002 | -34.404 | -17.3117
10.2063 | -34.4045 | -17.321
10.2124 | -34.4048 | -17.3304
10.2185 | -34.4053 | -17.3398
10.2246 | -34.4057 | -17.3491
10.2307 | -34.4061 | -17.3584
10.2368 | -34.4065 | -17.3674
10.2429 | -34.4069 | -17.3766
10.2491 | -34.4073 | -17.386
10.2552 | -34.4077 | -17.3953
10.2613 | -34.4081 | -17.4047
10.2674 | -34.4085 | -17.4139
10.2735 | -34.409 | -17.4234
10.2796 | -34.4094 | -17.4327
10.2857 | -34.4098 | -17.4421
10.2918 | -34.4102 | -17.4516
10.2979 | -34.4106 | -17.461
10.304 | -34.4111 | -17.4702
10.3101 | -34.4114 | -17.4796
10.3162 | -34.4118 | -17.489
10.3223 | -34.4122 | -17.4982
10.3284 | -34.4126 | -17.5074
10.3345 | -34.4131 | -17.5166
10.3406 | -34.4135 | -17.5257
10.3467 | -34.4139 | -17.535
10.3528 | -34.4143 | -17.5441
10.3589 | -34.4147 | -17.5534
10.365 | -34.4151 | -17.5628
10.3711 | -34.4155 | -17.5722
10.3772 | -34.4159 | -17.5814
10.3833 | -34.4163 | -17.5907
10.3895 | -34.4168 | -17.5998
10.3956 | -34.4172 | -17.6091
10.4017 | -34.4176 | -17.6182
10.4078 | -34.418 | -17.6274
10.4139 | -34.4184 | -17.6366
10.42 | -34.4188 | -17.646
10.4261 | -34.4193 | -17.6552
10.4322 | -34.4196 | -17.6644
10.4383 | -34.42 | -17.6737
10.4444 | -34.4204 | -17.6828
10.4505 | -34.4208 | -17.6922
10.4566 | -34.4212 | -17.7013
10.4627 | -34.4216 | -17.7105
10.4688 | -34.4221 | -17.7197
10.4749 | -34.4225 | -17.7289
10.481 | -34.4229 | -17.7382
10.4871 | -34.4233 | -17.7473
10.4932 | -34.4237 | -17.7564
10.4993 | -34.4241 | -17.7657
10.5054 | -34.4245 | -17.775
10.5115 | -34.4249 | -17.7843
10.5176 | -34.4254 | -17.7934
10.5237 | -34.4258 | -17.8028
10.5298 | -34.4262 | -17.8119
10.536 | -34.4266 | -17.821
10.5421 | -34.427 | -17.8302
10.5482 | -34.4274 | -17.8395
10.5543 | -34.4278 | -17.8486
10.5604 | -34.4283 | -17.8576
10.5665 | -34.4287 | -17.8667
10.5726 | -34.4291 | -17.8759
10.5787 | -34.4295 | -17.8851
10.5848 | -34.4299 | -17.8939
10.5909 | -34.4304 | -17.9032
10.597 | -34.4308 | -17.9124
10.6031 | -34.4312 | -17.9217
10.6092 | -34.4315 | -17.9309
10.6153 | -34.4319 | -17.9403
10.6214 | -34.4324 | -17.9494
10.6275 | -34.4328 | -17.9586
10.6336 | -34.4332 | -17.9679
10.6397 | -34.4336 | -17.977
10.6458 | -34.434 | -17.9861
10.6519 | -34.4344 | -17.9955
10.658 | -34.4348 | -18.0046
10.6641 | -34.4353 | -18.0138
10.6702 | -34.4357 | -18.0231
10.6764 | -34.4361 | -18.0323
10.6825 | -34.4365 | -18.0413
10.6886 | -34.4369 | -18.0503
10.6947 | -34.4373 | -18.0593
10.7008 | -34.4378 | -18.0685
10.7069 | -34.4382 | -18.0779
10.713 | -34.4386 | -18.0871
10.7191 | -34.439 | -18.0962
10.7252 | -34.4394 | -18.1051
10.7313 | -34.4398 | -18.1142
10.7374 | -34.4402 | -18.1234
10.7435 | -34.4406 | -18.1326
10.7496 | -34.441 | -18.1416
10.7557 | -34.4414 | -18.1509
10.7618 | -34.4419 | -18.1597
10.7679 | -34.4423 | -18.1688
10.774 | -34.4427 | -18.1781
10.7801 | -34.4431 | -18.1874
10.7862 | -34.4435 | -18.1963
10.7923 | -34.4439 | -18.2055
10.7984 | -34.4444 | -18.2146
10.8045 | -34.4448 | -18.2235
10.8106 | -34.4452 | -18.2327
10.8168 | -34.4456 | -18.2418
10.8229 | -34.446 | -18.2509
10.829 | -34.4464 | -18.26
10.8351 | -34.4469 | -18.2688
10.8412 | -34.4473 | -18.2778
10.8473 | -34.4477 | -18.2868
10.8534 | -34.4482 | -18.2959
10.8595 | -34.4486 | -18.305
10.8656 | -34.449 | -18.3138
10.8717 | -34.4494 | -18.3229
10.8778 | -34.4499 | -18.3321
10.8839 | -34.4503 | -18.3411
10.89 | -34.4507 | -18.3503
10.8961 | -34.4512 | -18.3594
10.9022 | -34.4516 | -18.3683
10.9083 | -34.452 | -18.3772
10.9144 | -34.4524 | -18.3863
10.9205 | -34.4528 | -18.3954
10.9266 | -34.4532 | -18.4044
10.9327 | -34.4537 | -18.4135
10.9388 | -34.4541 | -18.4224
10.9449 | -34.4545 | -18.4313
10.951 | -34.4549 | -18.4403
10.9571 | -34.4554 | -18.4494
10.9633 | -34.4558 | -18.4586
10.9694 | -34.4562 | -18.4679
10.9755 | -34.4566 | -18.4769
10.9816 | -34.457 | -18.4861
10.9877 | -34.4575 | -18.4951
10.9938 | -34.4579 | -18.5043
10.9999 | -34.4583 | -18.5132
11.006 | -34.4588 | -18.5221
11.0121 | -34.4592 | -18.5312
11.0182 | -34.4596 | -18.5403
11.0243 | -34.46 | -18.5494
11.0304 | -34.4604 | -18.5584
11.0365 | -34.4609 | -18.5674
11.0426 | -34.4613 | -18.5764
11.0487 | -34.4618 | -18.5855
11.0548 | -34.4622 | -18.5943
11.0609 | -34.4626 | -18.6035
11.067 | -34.4631 | -18.6125
11.0731 | -34.4635 | -18.6214
11.0792 | -34.4639 | -18.6304
11.0853 | -34.4644 | -18.6395
11.0914 | -34.4648 | -18.6486
11.0975 | -34.4652 | -18.6576
11.1037 | -34.4656 | -18.6667
11.1098 | -34.4661 | -18.6759
11.1159 | -34.4665 | -18.6848
11.122 | -34.4669 | -18.6937
11.1281 | -34.4673 | -18.7028
11.1342 | -34.4678 | -18.7119
11.1403 | -34.4682 | -18.721
11.1464 | -34.4686 | -18.7301
11.1525 | -34.4691 | -18.7388
11.1586 | -34.4695 | -18.748
11.1647 | -34.4699 | -18.7569
11.1708 | -34.4703 | -18.7659
11.1769 | -34.4708 | -18.775
11.183 | -34.4712 | -18.7837
11.1891 | -34.4716 | -18.7926
11.1952 | -34.4721 | -18.8016
11.2013 | -34.4725 | -18.8107
11.2074 | -34.473 | -18.8197
11.2135 | -34.4734 | -18.8289
11.2196 | -34.4738 | -18.8379
11.2257 | -34.4743 | -18.8471
11.2318 | -34.4747 | -18.8561
11.2379 | -34.4752 | -18.8652
11.244 | -34.4756 | -18.8743
11.2502 | -34.4761 | -18.8833
11.2563 | -34.4765 | -18.8922
11.2624 | -34.4769 | -18.9012
11.2685 | -34.4774 | -18.9101
11.2746 | -34.4778 | -18.9191
11.2807 | -34.4782 | -18.9282
11.2868 | -34.4787 | -18.9373
11.2929 | -34.4791 | -18.946
11.299 | -34.4796 | -18.955
11.3051 | -34.48 | -18.9641
11.3112 | -34.4804 | -18.9731
11.3173 | -34.4808 | -18.9818
11.3234 | -34.4813 | -18.9906
11.3295 | -34.4817 | -18.9996
11.3356 | -34.4822 | -19.0087
11.3417 | -34.4826 | -19.0176
11.3478 | -34.483 | -19.0268
11.3539 | -34.4834 | -19.0357
11.36 | -34.4839 | -19.0447
11.3661 | -34.4843 | -19.0535
11.3722 | -34.4848 | -19.0622
11.3783 | -34.4852 | -19.0712
11.3844 | -34.4857 | -19.0801
11.3906 | -34.4862 | -19.089
11.3967 | -34.4866 | -19.0979
11.4028 | -34.4871 | -19.107
11.4089 | -34.4875 | -19.1159
11.415 | -34.4879 | -19.1248
11.4211 | -34.4883 | -19.1339
11.4272 | -34.4888 | -19.1429
11.4333 | -34.4892 | -19.1517
11.4394 | -34.4896 | -19.1605
11.4455 | -34.4901 | -19.1694
11.4516 | -34.4905 | -19.1783
11.4577 | -34.4909 | -19.1872
11.4638 | -34.4914 | -19.1961
11.4699 | -34.4918 | -19.2052
11.476 | -34.4923 | -19.2143
11.4821 | -34.4927 | -19.2232
11.4882 | -34.4932 | -19.2324
11.4943 | -34.4936 | -19.2415
11.5004 | -34.4941 | -19.2501
11.5065 | -34.4945 | -19.259
11.5126 | -34.495 | -19.2678
11.5187 | -34.4954 | -19.2769
11.5248 | -34.4958 | -19.286
11.5309 | -34.4963 | -19.2949
11.5371 | -34.4967 | -19.3037
11.5432 | -34.4972 | -19.3127
11.5493 | -34.4976 | -19.3218
11.5554 | -34.498 | -19.3307
11.5615 | -34.4985 | -19.3398
11.5676 | -34.499 | -19.3488
11.5737 | -34.4994 | -19.3577
11.5798 | -34.4999 | -19.3666
11.5859 | -34.5004 | -19.3755
11.592 | -34.5008 | -19.3844
11.5981 | -34.5012 | -19.3933
11.6042 | -34.5017 | -19.4023
11.6103 | -34.5021 | -19.4111
11.6164 | -34.5026 | -19.4198
11.6225 | -34.503 | -19.4287
11.6286 | -34.5034 | -19.4376
11.6347 | -34.5039 | -19.4466
11.6408 | -34.5043 | -19.4556
11.6469 | -34.5048 | -19.4644
11.653 | -34.5052 | -19.4732
11.6591 | -34.5056 | -19.4822
11.6652 | -34.5061 | -19.491
11.6713 | -34.5065 | -19.4998
11.6775 | -34.507 | -19.5085
11.6836 | -34.5074 | -19.5174
11.6897 | -34.5079 | -19.5264
11.6958 | -34.5083 | -19.5354
11.7019 | -34.5088 | -19.5443
11.708 | -34.5092 | -19.5531
11.7141 | -34.5097 | -19.5618
11.7202 | -34.5102 | -19.5707
11.7263 | -34.5106 | -19.5795
11.7324 | -34.511 | -19.5886
11.7385 | -34.5115 | -19.5976
11.7446 | -34.5119 | -19.6065
11.7507 | -34.5124 | -19.6154
11.7568 | -34.5128 | -19.6243
11.7629 | -34.5133 | -19.6333
11.769 | -34.5137 | -19.6421
11.7751 | -34.5141 | -19.651
11.7812 | -34.5146 | -19.6598
11.7873 | -34.5151 | -19.6687
11.7934 | -34.5155 | -19.6777
11.7995 | -34.516 | -19.6866
11.8056 | -34.5164 | -19.6953
11.8117 | -34.5169 | -19.7042
11.8178 | -34.5173 | -19.7129
11.824 | -34.5178 | -19.7217
11.8301 | -34.5183 | -19.7305
11.8362 | -34.5187 | -19.7392
11.8423 | -34.5191 | -19.748
11.8484 | -34.5196 | -19.7569
11.8545 | -34.5201 | -19.7658
11.8606 | -34.5205 | -19.7746
11.8667 | -34.521 | -19.7836
11.8728 | -34.5214 | -19.7922
11.8789 | -34.5218 | -19.801
11.885 | -34.5223 | -19.8099
11.8911 | -34.5228 | -19.8186
11.8972 | -34.5232 | -19.8274
11.9033 | -34.5237 | -19.8362
11.9094 | -34.5241 | -19.8449
11.9155 | -34.5246 | -19.8537
11.9216 | -34.5251 | -19.8625
11.9277 | -34.5255 | -19.8712
11.9338 | -34.5259 | -19.8801
11.9399 | -34.5264 | -19.8889
11.946 | -34.5269 | -19.8977
11.9521 | -34.5273 | -19.9064
11.9582 | -34.5278 | -19.9151
11.9644 | -34.5283 | -19.9238
11.9705 | -34.5287 | -19.9326
11.9766 | -34.5292 | -19.9413
11.9827 | -34.5297 | -19.9502
11.9888 | -34.5302 | -19.959
11.9949 | -34.5306 | -19.9677
12.001 | -34.5311 | -19.9765
12.0071 | -34.5315 | -19.9852
12.0132 | -34.532 | -19.9938
12.0193 | -34.5324 | -20.0026
12.0254 | -34.5328 | -20.0115
12.0315 | -34.5333 | -20.02
12.0376 | -34.5337 | -20.0287
12.0437 | -34.5342 | -20.0373
12.0498 | -34.5346 | -20.0463
12.0559 | -34.5351 | -20.0549
12.062 | -34.5356 | -20.0633
12.0681 | -34.536 | -20.0722
12.0742 | -34.5365 | -20.081
12.0803 | -34.537 | -20.09
12.0864 | -34.5374 | -20.0987
12.0925 | -34.5379 | -20.1076
12.0986 | -34.5383 | -20.1164
12.1047 | -34.5388 | -20.1253
12.1109 | -34.5393 | -20.1339
12.117 | -34.5397 | -20.1425
12.1231 | -34.5402 | -20.1512
12.1292 | -34.5406 | -20.16
12.1353 | -34.5411 | -20.1688
12.1414 | -34.5416 | -20.1777
12.1475 | -34.542 | -20.1865
12.1536 | -34.5424 | -20.1953
12.1597 | -34.5429 | -20.2042
12.1658 | -34.5434 | -20.2129
12.1719 | -34.5439 | -20.2216
12.178 | -34.5443 | -20.2304
12.1841 | -34.5448 | -20.2392
12.1902 | -34.5452 | -20.2479
12.1963 | -34.5457 | -20.2567
12.2024 | -34.5462 | -20.2652
12.2085 | -34.5466 | -20.2739
12.2146 | -34.5471 | -20.2827
12.2207 | -34.5475 | -20.2914
12.2268 | -34.548 | -20.3001
12.2329 | -34.5485 | -20.3087
12.239 | -34.549 | -20.3172
12.2451 | -34.5494 | -20.3259
12.2513 | -34.5499 | -20.3348
12.2574 | -34.5503 | -20.3435
12.2635 | -34.5508 | -20.3522
12.2696 | -34.5512 | -20.3608
12.2757 | -34.5517 | -20.3692
12.2818 | -34.5522 | -20.378
12.2879 | -34.5526 | -20.3867
12.294 | -34.5531 | -20.3953
12.3001 | -34.5535 | -20.404
12.3062 | -34.554 | -20.4128
12.3123 | -34.5545 | -20.4213
12.3184 | -34.555 | -20.4299
12.3245 | -34.5554 | -20.4384
12.3306 | -34.5559 | -20.4471
12.3367 | -34.5564 | -20.4557
12.3428 | -34.5568 | -20.4644
12.3489 | -34.5573 | -20.4728
12.355 | -34.5577 | -20.4816
12.3611 | -34.5582 | -20.4904
12.3672 | -34.5586 | -20.499
12.3733 | -34.5591 | -20.5077
12.3794 | -34.5596 | -20.516
12.3855 | -34.56 | -20.5246
12.3916 | -34.5605 | -20.5333
12.3978 | -34.561 | -20.5421
12.4039 | -34.5614 | -20.5508
12.41 | -34.5619 | -20.5595
12.4161 | -34.5624 | -20.5678
12.4222 | -34.5628 | -20.5765
12.4283 | -34.5633 | -20.5851
12.4344 | -34.5638 | -20.5938
12.4405 | -34.5642 | -20.6025
12.4466 | -34.5646 | -20.611
12.4527 | -34.5651 | -20.6197
12.4588 | -34.5656 | -20.6283
12.4649 | -34.566 | -20.637
12.471 | -34.5665 | -20.6457
12.4771 | -34.5669 | -20.6543
12.4832 | -34.5674 | -20.6629
12.4893 | -34.5678 | -20.6715
12.4954 | -34.5683 | -20.6799
12.5015 | -34.5688 | -20.6885
12.5076 | -34.5692 | -20.6973
12.5137 | -34.5697 | -20.7059
12.5198 | -34.5702 | -20.7145
12.5259 | -34.5706 | -20.7233
12.532 | -34.571 | -20.7318
12.5382 | -34.5715 | -20.7404
12.5443 | -34.572 | -20.749
12.5504 | -34.5724 | -20.7577
12.5565 | -34.5729 | -20.7661
12.5626 | -34.5734 | -20.7749
12.5687 | -34.5738 | -20.7835
12.5748 | -34.5743 | -20.7921
12.5809 | -34.5747 | -20.8007
12.587 | -34.5752 | -20.8095
12.5931 | -34.5757 | -20.8178
12.5992 | -34.5762 | -20.8264
12.6053 | -34.5766 | -20.835
12.6114 | -34.5771 | -20.8436
12.6175 | -34.5776 | -20.8524
12.6236 | -34.578 | -20.861
12.6297 | -34.5785 | -20.8696
12.6358 | -34.5789 | -20.878
12.6419 | -34.5794 | -20.8864
12.648 | -34.5799 | -20.8952
12.6541 | -34.5803 | -20.9037
12.6602 | -34.5808 | -20.9119
12.6663 | -34.5812 | -20.9205
12.6724 | -34.5817 | -20.929
12.6785 | -34.5822 | -20.9376
12.6847 | -34.5826 | -20.946
12.6908 | -34.5831 | -20.9546
12.6969 | -34.5836 | -20.963
12.703 | -34.584 | -20.9716
12.7091 | -34.5845 | -20.9802
12.7152 | -34.5849 | -20.9888
12.7213 | -34.5854 | -20.9974
12.7274 | -34.5858 | -21.0059
12.7335 | -34.5863 | -21.0141
12.7396 | -34.5868 | -21.0224
12.7457 | -34.5872 | -21.0311
12.7518 | -34.5877 | -21.0396
12.7579 | -34.5881 | -21.0482
12.764 | -34.5886 | -21.0569
12.7701 | -34.589 | -21.0655
12.7762 | -34.5895 | -21.0739
12.7823 | -34.59 | -21.0825
12.7884 | -34.5905 | -21.0911
12.7945 | -34.5909 | -21.0998
12.8006 | -34.5914 | -21.1082
12.8067 | -34.5918 | -21.1167
12.8128 | -34.5922 | -21.125
12.8189 | -34.5927 | -21.1335
12.8251 | -34.5931 | -21.1421
12.8312 | -34.5936 | -21.1508
12.8373 | -34.5941 | -21.1591
12.8434 | -34.5945 | -21.1678
12.8495 | -34.595 | -21.1762
12.8556 | -34.5954 | -21.1847
12.8617 | -34.5959 | -21.1932
12.8678 | -34.5963 | -21.2019
12.8739 | -34.5968 | -21.2103
12.88 | -34.5972 | -21.2187
12.8861 | -34.5977 | -21.2272
12.8922 | -34.5982 | -21.2358
12.8983 | -34.5986 | -21.2443
12.9044 | -34.599 | -21.2526
12.9105 | -34.5995 | -21.2609
12.9166 | -34.6 | -21.2693
12.9227 | -34.6004 | -21.2779
12.9288 | -34.6009 | -21.2864
12.9349 | -34.6013 | -21.295
12.941 | -34.6018 | -21.3035
12.9471 | -34.6023 | -21.312
12.9532 | -34.6027 | -21.3206
12.9593 | -34.6032 | -21.3293
12.9654 | -34.6037 | -21.3379
12.9716 | -34.6041 | -21.3465
12.9777 | -34.6046 | -21.3551
12.9838 | -34.6051 | -21.3635
12.9899 | -34.6055 | -21.3721
12.996 | -34.6059 | -21.3806
13.0021 | -34.6064 | -21.3891
13.0082 | -34.6069 | -21.3974
13.0143 | -34.6074 | -21.4056
13.0204 | -34.6078 | -21.4139
13.0265 | -34.6083 | -21.4223
13.0326 | -34.6087 | -21.4307
13.0387 | -34.6092 | -21.4393
13.0448 | -34.6096 | -21.4478
13.0509 | -34.6101 | -21.4562
13.057 | -34.6106 | -21.4646
13.0631 | -34.611 | -21.473
13.0692 | -34.6115 | -21.4813
13.0753 | -34.612 | -21.4899
13.0814 | -34.6124 | -21.4985
13.0875 | -34.6129 | -21.5068
13.0936 | -34.6133 | -21.5154
13.0997 | -34.6137 | -21.5237
13.1058 | -34.6142 | -21.5321
13.112 | -34.6146 | -21.5406
13.1181 | -34.6151 | -21.5488
13.1242 | -34.6155 | -21.5573
13.1303 | -34.616 | -21.5657
13.1364 | -34.6164 | -21.5742
13.1425 | -34.6169 | -21.5825
13.1486 | -34.6173 | -21.591
13.1547 | -34.6178 | -21.5995
13.1608 | -34.6183 | -21.6079
13.1669 | -34.6187 | -21.6161
13.173 | -34.6192 | -21.6247
13.1791 | -34.6196 | -21.6331
13.1852 | -34.6201 | -21.6415
13.1913 | -34.6205 | -21.6498
13.1974 | -34.621 | -21.6583
13.2035 | -34.6214 | -21.6667
13.2096 | -34.6219 | -21.675
13.2157 | -34.6223 | -21.6835
13.2218 | -34.6228 | -21.6919
13.2279 | -34.6233 | -21.7
13.234 | -34.6237 | -21.7084
13.2401 | -34.6242 | -21.7168
13.2462 | -34.6246 | -21.7251
13.2524 | -34.6251 | -21.7335
13.2585 | -34.6256 | -21.7419
13.2646 | -34.626 | -21.7502
13.2707 | -34.6265 | -21.7586
13.2768 | -34.627 | -21.7673
13.2829 | -34.6274 | -21.7757
13.289 | -34.6279 | -21.7839
13.2951 | -34.6284 | -21.7922
13.3012 | -34.6288 | -21.8006
13.3073 | -34.6293 | -21.8091
13.3134 | -34.6297 | -21.8175
13.3195 | -34.6302 | -21.8259
13.3256 | -34.6306 | -21.8342
13.3317 | -34.6311 | -21.8424
13.3378 | -34.6315 | -21.8509
13.3439 | -34.632 | -21.8594
13.35 | -34.6324 | -21.8679
13.3561 | -34.6329 | -21.8761
13.3622 | -34.6333 | -21.8845
13.3683 | -34.6338 | -21.8927
13.3744 | -34.6342 | -21.9009
13.3805 | -34.6347 | -21.9094
13.3866 | -34.6351 | -21.9177
13.3927 | -34.6355 | -21.9262
13.3989 | -34.636 | -21.9344
13.405 | -34.6364 | -21.9428
13.4111 | -34.6369 | -21.9511
13.4172 | -34.6373 | -21.9595
13.4233 | -34.6378 | -21.968
13.4294 | -34.6382 | -21.9765
13.4355 | -34.6387 | -21.9846
13.4416 | -34.6391 | -21.9928
13.4477 | -34.6396 | -22.0012
13.4538 | -34.64 | -22.0096
13.4599 | -34.6405 | -22.0181
13.466 | -34.6409 | -22.0265
13.4721 | -34.6414 | -22.0346
13.4782 | -34.6419 | -22.0429
13.4843 | -34.6423 | -22.0514
13.4904 | -34.6428 | -22.0597
13.4965 | -34.6432 | -22.0681
13.5026 | -34.6437 | -22.0765
13.5087 | -34.6441 | -22.0847
13.5148 | -34.6446 | -22.0931
13.5209 | -34.645 | -22.1017
13.527 | -34.6454 | -22.11
13.5331 | -34.6459 | -22.1182
13.5393 | -34.6464 | -22.1264
13.5454 | -34.6469 | -22.1345
13.5515 | -34.6473 | -22.1431
13.5576 | -34.6478 | -22.1515
13.5637 | -34.6482 | -22.1598
13.5698 | -34.6487 | -22.1682
13.5759 | -34.6491 | -22.1762
13.582 | -34.6496 | -22.1844
13.5881 | -34.65 | -22.1926
13.5942 | -34.6505 | -22.2011
13.6003 | -34.651 | -22.2094
13.6064 | -34.6514 | -22.2181
13.6125 | -34.6519 | -22.2265
13.6186 | -34.6524 | -22.2348
13.6247 | -34.6528 | -22.243
13.6308 | -34.6533 | -22.2513
13.6369 | -34.6537 | -22.2598
13.643 | -34.6542 | -22.2683
13.6491 | -34.6547 | -22.2765
13.6552 | -34.6551 | -22.2848
13.6613 | -34.6556 | -22.2932
13.6674 | -34.656 | -22.3015
13.6735 | -34.6565 | -22.3098
13.6796 | -34.657 | -22.3182
13.6858 | -34.6574 | -22.3265
13.6919 | -34.6579 | -22.3349
13.698 | -34.6583 | -22.3432
13.7041 | -34.6588 | -22.3516
13.7102 | -34.6592 | -22.3599
13.7163 | -34.6597 | -22.3682
13.7224 | -34.6602 | -22.3765
13.7285 | -34.6607 | -22.3848
13.7346 | -34.6611 | -22.3931
13.7407 | -34.6616 | -22.4015
13.7468 | -34.662 | -22.4099
13.7529 | -34.6625 | -22.4181
13.759 | -34.6629 | -22.4264
13.7651 | -34.6634 | -22.4346
13.7712 | -34.6638 | -22.443
13.7773 | -34.6643 | -22.4513
13.7834 | -34.6647 | -22.4595
13.7895 | -34.6652 | -22.4678
13.7956 | -34.6656 | -22.4762
13.8017 | -34.6661 | -22.4844
13.8078 | -34.6665 | -22.4928
13.8139 | -34.667 | -22.5011
13.82 | -34.6674 | -22.5093
13.8262 | -34.6679 | -22.5175
13.8323 | -34.6683 | -22.526
13.8384 | -34.6688 | -22.5342
13.8445 | -34.6692 | -22.5425
13.8506 | -34.6697 | -22.5507
13.8567 | -34.6701 | -22.5589
13.8628 | -34.6706 | -22.5673
13.8689 | -34.671 | -22.5757
13.875 | -34.6715 | -22.5838
13.8811 | -34.6719 | -22.5921
13.8872 | -34.6724 | -22.6003
13.8933 | -34.6728 | -22.6084
13.8994 | -34.6733 | -22.6167
13.9055 | -34.6737 | -22.6252
13.9116 | -34.6742 | -22.6334
13.9177 | -34.6747 | -22.6418
13.9238 | -34.6751 | -22.65
13.9299 | -34.6756 | -22.6582
13.936 | -34.676 | -22.6665
13.9421 | -34.6764 | -22.6746
13.9482 | -34.6769 | -22.6828
13.9543 | -34.6773 | -22.6913
13.9604 | -34.6778 | -22.6996
13.9665 | -34.6783 | -22.7077
13.9727 | -34.6787 | -22.7161
13.9788 | -34.6792 | -22.7243
13.9849 | -34.6797 | -22.7326
13.991 | -34.6801 | -22.7409
13.9971 | -34.6806 | -22.7492
14.0032 | -34.681 | -22.7575
14.0093 | -34.6815 | -22.7657
14.0154 | -34.6819 | -22.7739
14.0215 | -34.6824 | -22.782
14.0276 | -34.6828 | -22.7902
14.0337 | -34.6833 | -22.7985
14.0398 | -34.6838 | -22.8066
14.0459 | -34.6843 | -22.8149
14.052 | -34.6847 | -22.8233
14.0581 | -34.6852 | -22.8316
14.0642 | -34.6856 | -22.8399
14.0703 | -34.6861 | -22.8481
14.0764 | -34.6865 | -22.8562
14.0825 | -34.687 | -22.8645
14.0886 | -34.6874 | -22.8727
14.0947 | -34.6879 | -22.881
14.1008 | -34.6883 | -22.8892
14.1069 | -34.6888 | -22.8974
14.1131 | -34.6893 | -22.9059
14.1192 | -34.6897 | -22.9142
14.1253 | -34.6902 | -22.9222
14.1314 | -34.6906 | -22.9305
14.1375 | -34.6911 | -22.9386
14.1436 | -34.6916 | -22.9468
14.1497 | -34.692 | -22.955
14.1558 | -34.6925 | -22.9633
14.1619 | -34.6929 | -22.9715
14.168 | -34.6934 | -22.9798
14.1741 | -34.6939 | -22.9878
14.1802 | -34.6943 | -22.996
14.1863 | -34.6948 | -23.0042
14.1924 | -34.6953 | -23.0124
14.1985 | -34.6958 | -23.0206
14.2046 | -34.6962 | -23.0289
14.2107 | -34.6967 | -23.0368
14.2168 | -34.6971 | -23.0451
14.2229 | -34.6976 | -23.0533
14.229 | -34.698 | -23.0616
14.2351 | -34.6985 | -23.0698
14.2412 | -34.6989 | -23.078
14.2473 | -34.6994 | -23.0861
14.2534 | -34.6999 | -23.0943
14.2596 | -34.7003 | -23.1024
14.2657 | -34.7007 | -23.1106
14.2718 | -34.7012 | -23.1188
14.2779 | -34.7016 | -23.1271
14.284 | -34.7021 | -23.1354
14.2901 | -34.7026 | -23.1436
14.2962 | -34.7031 | -23.1519
14.3023 | -34.7035 | -23.1601
14.3084 | -34.704 | -23.1683
14.3145 | -34.7045 | -23.1765
14.3206 | -34.7049 | -23.1849
14.3267 | -34.7054 | -23.1931
14.3328 | -34.7059 | -23.2013
14.3389 | -34.7063 | -23.2095
14.345 | -34.7068 | -23.2176
14.3511 | -34.7073 | -23.2258
14.3572 | -34.7077 | -23.234
14.3633 | -34.7082 | -23.2423
14.3694 | -34.7087 | -23.2505
14.3755 | -34.7091 | -23.2588
14.3816 | -34.7096 | -23.2671
14.3877 | -34.7101 | -23.2752
14.3938 | -34.7106 | -23.2836
14.4 | -34.7111 | -23.2917
14.4061 | -34.7115 | -23.2999
14.4122 | -34.712 | -23.3081
14.4183 | -34.7124 | -23.3162
14.4244 | -34.7129 | -23.3245
14.4305 | -34.7134 | -23.3325
14.4366 | -34.7139 | -23.3407
14.4427 | -34.7143 | -23.3492
14.4488 | -34.7148 | -23.3575
14.4549 | -34.7153 | -23.3658
14.461 | -34.7157 | -23.3738
14.4671 | -34.7162 | -23.382
14.4732 | -34.7167 | -23.3903
14.4793 | -34.7171 | -23.3985
14.4854 | -34.7176 | -23.4066
14.4915 | -34.7181 | -23.4145
14.4976 | -34.7186 | -23.4225
14.5037 | -34.7191 | -23.4305
14.5098 | -34.7195 | -23.4388
14.5159 | -34.72 | -23.4471
14.522 | -34.7204 | -23.4552
14.5281 | -34.7209 | -23.4631
14.5342 | -34.7214 | -23.4713
14.5403 | -34.7219 | -23.4795
14.5465 | -34.7224 | -23.4876
14.5526 | -34.7228 | -23.496
14.5587 | -34.7233 | -23.5042
14.5648 | -34.7238 | -23.5123
14.5709 | -34.7243 | -23.5204
14.577 | -34.7248 | -23.5285
14.5831 | -34.7252 | -23.5366
14.5892 | -34.7257 | -23.5448
14.5953 | -34.7262 | -23.5528
14.6014 | -34.7267 | -23.561
14.6075 | -34.7271 | -23.569
14.6136 | -34.7275 | -23.5771
14.6197 | -34.728 | -23.5853
14.6258 | -34.7285 | -23.5935
14.6319 | -34.729 | -23.6014
14.638 | -34.7294 | -23.6096
14.6441 | -34.7299 | -23.6179
14.6502 | -34.7304 | -23.6262
14.6563 | -34.7309 | -23.6343
14.6624 | -34.7314 | -23.6424
14.6685 | -34.7319 | -23.6503
14.6746 | -34.7323 | -23.6585
14.6807 | -34.7328 | -23.6666
14.6869 | -34.7333 | -23.6747
14.693 | -34.7338 | -23.6828
14.6991 | -34.7342 | -23.6909
14.7052 | -34.7348 | -23.6988
14.7113 | -34.7352 | -23.7071
14.7174 | -34.7357 | -23.7153
14.7235 | -34.7362 | -23.7234
14.7296 | -34.7366 | -23.7315
14.7357 | -34.7371 | -23.7395
14.7418 | -34.7376 | -23.7475
14.7479 | -34.7381 | -23.7555
14.754 | -34.7386 | -23.7637
14.7601 | -34.7391 | -23.7717
14.7662 | -34.7396 | -23.7798
14.7723 | -34.74 | -23.7877
14.7784 | -34.7405 | -23.7959
14.7845 | -34.741 | -23.8042
14.7906 | -34.7414 | -23.8123
14.7967 | -34.7419 | -23.8205
14.8028 | -34.7424 | -23.8285
14.8089 | -34.7429 | -23.8365
14.815 | -34.7434 | -23.8447
14.8211 | -34.7438 | -23.8527
14.8272 | -34.7443 | -23.8609
14.8334 | -34.7448 | -23.8689
14.8395 | -34.7453 | -23.8769
14.8456 | -34.7458 | -23.8849
14.8517 | -34.7462 | -23.8931
14.8578 | -34.7467 | -23.9012
14.8639 | -34.7472 | -23.9093
14.87 | -34.7477 | -23.9173
14.8761 | -34.7482 | -23.9252
14.8822 | -34.7487 | -23.9332
14.8883 | -34.7491 | -23.9414
14.8944 | -34.7496 | -23.9493
14.9005 | -34.7501 | -23.9574
14.9066 | -34.7506 | -23.9654
14.9127 | -34.7511 | -23.9735
14.9188 | -34.7515 | -23.9816
14.9249 | -34.752 | -23.9897
14.931 | -34.7525 | -23.9978
14.9371 | -34.753 | -24.0058
14.9432 | -34.7535 | -24.0139
14.9493 | -34.7539 | -24.022
14.9554 | -34.7544 | -24.0301
14.9615 | -34.7549 | -24.0382
14.9676 | -34.7554 | -24.0465
14.9738 | -34.7558 | -24.0545
14.9799 | -34.7563 | -24.0628
14.986 | -34.7568 | -24.0708
14.9921 | -34.7573 | -24.0789
14.9982 | -34.7577 | -24.087
15.0043 | -34.7582 | -24.095
15.0104 | -34.7587 | -24.1032
15.0165 | -34.7592 | -24.1114
15.0226 | -34.7597 | -24.1192
15.0287 | -34.7602 | -24.1273
15.0348 | -34.7607 | -24.1355
15.0409 | -34.7612 | -24.1435
15.047 | -34.7617 | -24.1518
15.0531 | -34.7622 | -24.1596
15.0592 | -34.7626 | -24.1678
15.0653 | -34.7631 | -24.1758
15.0714 | -34.7636 | -24.1839
15.0775 | -34.7641 | -24.1918
15.0836 | -34.7646 | -24.1998
15.0897 | -34.7651 | -24.2077
15.0958 | -34.7656 | -24.2157
15.1019 | -34.7661 | -24.2237
15.108 | -34.7666 | -24.2316
15.1141 | -34.7671 | -24.2396
15.1203 | -34.7676 | -24.2477
15.1264 | -34.7681 | -24.2556
15.1325 | -34.7685 | -24.2638
15.1386 | -34.769 | -24.2718
15.1447 | -34.7695 | -24.2798
15.1508 | -34.77 | -24.2879
15.1569 | -34.7705 | -24.296
15.163 | -34.771 | -24.3039
15.1691 | -34.7715 | -24.3119
15.1752 | -34.772 | -24.3199
15.1813 | -34.7725 | -24.3279
15.1874 | -34.773 | -24.336
15.1935 | -34.7735 | -24.3439
15.1996 | -34.774 | -24.3519
15.2057 | -34.7744 | -24.3598
15.2118 | -34.7749 | -24.3678
15.2179 | -34.7754 | -24.3757
15.224 | -34.7759 | -24.3837
15.2301 | -34.7764 | -24.3915
15.2362 | -34.7769 | -24.3996
15.2423 | -34.7774 | -24.4077
15.2484 | -34.7779 | -24.4158
15.2545 | -34.7784 | -24.4239
15.2607 | -34.7789 | -24.432
15.2668 | -34.7794 | -24.4398
15.2729 | -34.7799 | -24.4478
15.279 | -34.7804 | -24.4559
15.2851 | -34.7809 | -24.4638
15.2912 | -34.7814 | -24.4719
15.2973 | -34.7818 | -24.4801
15.3034 | -34.7824 | -24.4879
15.3095 | -34.7829 | -24.4959
15.3156 | -34.7834 | -24.5038
15.3217 | -34.7839 | -24.5117
15.3278 | -34.7844 | -24.5198
15.3339 | -34.7849 | -24.5276
15.34 | -34.7854 | -24.5354
15.3461 | -34.7859 | -24.5434
15.3522 | -34.7864 | -24.5514
15.3583 | -34.7868 | -24.5597
15.3644 | -34.7873 | -24.5676
15.3705 | -34.7878 | -24.5752
15.3766 | -34.7883 | -24.5833
15.3827 | -34.7888 | -24.5912
15.3888 | -34.7893 | -24.5992
15.3949 | -34.7898 | -24.6072
15.401 | -34.7903 | -24.6153
15.4072 | -34.7908 | -24.623
15.4133 | -34.7913 | -24.6311
15.4194 | -34.7918 | -24.639
15.4255 | -34.7923 | -24.6469
15.4316 | -34.7927 | -24.6549
15.4377 | -34.7932 | -24.663
15.4438 | -34.7938 | -24.6708
15.4499 | -34.7943 | -24.6787
15.456 | -34.7948 | -24.6867
15.4621 | -34.7953 | -24.6945
15.4682 | -34.7958 | -24.7026
15.4743 | -34.7963 | -24.7104
15.4804 | -34.7967 | -24.7184
15.4865 | -34.7972 | -24.7264
15.4926 | -34.7977 | -24.7345
15.4987 | -34.7982 | -24.7424
15.5048 | -34.7987 | -24.7505
15.5109 | -34.7992 | -24.7585
15.517 | -34.7998 | -24.7665
15.5231 | -34.8003 | -24.7744
15.5292 | -34.8007 | -24.7823
15.5353 | -34.8012 | -24.7904
15.5414 | -34.8017 | -24.7984
15.5476 | -34.8022 | -24.806
15.5537 | -34.8027 | -24.8141
15.5598 | -34.8032 | -24.8222
15.5659 | -34.8037 | -24.8301
15.572 | -34.8042 | -24.8381
15.5781 | -34.8047 | -24.8459
15.5842 | -34.8052 | -24.8537
15.5903 | -34.8058 | -24.8616
15.5964 | -34.8063 | -24.8694
15.6025 | -34.8068 | -24.8773
15.6086 | -34.8073 | -24.8852
15.6147 | -34.8078 | -24.8932
15.6208 | -34.8083 | -24.9008
15.6269 | -34.8088 | -24.9086
15.633 | -34.8093 | -24.9166
15.6391 | -34.8098 | -24.9246
15.6452 | -34.8103 | -24.9326
15.6513 | -34.8108 | -24.9402
15.6574 | -34.8113 | -24.9482
15.6635 | -34.8118 | -24.956
15.6696 | -34.8123 | -24.9639
15.6757 | -34.8128 | -24.972
15.6818 | -34.8133 | -24.9799
15.688 | -34.8138 | -24.9877
15.6941 | -34.8143 | -24.9956
15.7002 | -34.8148 | -25.0034
15.7063 | -34.8153 | -25.0113
15.7124 | -34.8158 | -25.0194
15.7185 | -34.8162 | -25.0272
15.7246 | -34.8168 | -25.0348
15.7307 | -34.8173 | -25.0426
15.7368 | -34.8178 | -25.0507
15.7429 | -34.8183 | -25.0583
15.749 | -34.8188 | -25.0661
15.7551 | -34.8193 | -25.0741
15.7612 | -34.8198 | -25.0818
15.7673 | -34.8203 | -25.0897
15.7734 | -34.8208 | -25.0975
15.7795 | -34.8213 | -25.1053
15.7856 | -34.8218 | -25.1132
15.7917 | -34.8223 | -25.1209
15.7978 | -34.8228 | -25.1289
15.8039 | -34.8233 | -25.1368
15.81 | -34.8238 | -25.1447
15.8161 | -34.8243 | -25.1525
15.8222 | -34.8249 | -25.1604
15.8283 | -34.8254 | -25.168
15.8345 | -34.8259 | -25.1759
15.8406 | -34.8264 | -25.1836
15.8467 | -34.8269 | -25.1915
15.8528 | -34.8274 | -25.1993
15.8589 | -34.8279 | -25.2071
15.865 | -34.8285 | -25.2147
15.8711 | -34.829 | -25.2226
15.8772 | -34.8295 | -25.2303
15.8833 | -34.83 | -25.2382
15.8894 | -34.8305 | -25.2462
15.8955 | -34.8309 | -25.2541
15.9016 | -34.8315 | -25.2617
15.9077 | -34.832 | -25.2697
15.9138 | -34.8324 | -25.2775
15.9199 | -34.8329 | -25.2851
15.926 | -34.8334 | -25.2928
15.9321 | -34.834 | -25.3004
15.9382 | -34.8344 | -25.3082
15.9443 | -34.835 | -25.316
15.9504 | -34.8355 | -25.3238
15.9565 | -34.836 | -25.3315
15.9626 | -34.8365 | -25.3394
15.9687 | -34.837 | -25.347
15.9749 | -34.8375 | -25.3547
15.981 | -34.838 | -25.3625
15.9871 | -34.8385 | -25.3703
15.9932 | -34.839 | -25.378
15.9993 | -34.8395 | -25.3859
16.0054 | -34.84 | -25.3934
16.0115 | -34.8405 | -25.4011
16.0176 | -34.841 | -25.4091
16.0237 | -34.8415 | -25.4169
16.0298 | -34.842 | -25.4246
16.0359 | -34.8425 | -25.4325
16.042 | -34.8431 | -25.4402
16.0481 | -34.8435 | -25.448
16.0542 | -34.8441 | -25.4559
16.0603 | -34.8446 | -25.4637
16.0664 | -34.8451 | -25.4715
16.0725 | -34.8456 | -25.479
16.0786 | -34.8461 | -25.4868
16.0847 | -34.8466 | -25.4947
16.0908 | -34.8471 | -25.5024
16.0969 | -34.8476 | -25.5104
16.103 | -34.8481 | -25.5183
16.1091 | -34.8486 | -25.5259
16.1152 | -34.8491 | -25.5336
16.1214 | -34.8496 | -25.5413
16.1275 | -34.8501 | -25.5491
16.1336 | -34.8506 | -25.5569
16.1397 | -34.8511 | -25.5646
16.1458 | -34.8516 | -25.5724
16.1519 | -34.8522 | -25.5802
16.158 | -34.8527 | -25.5882
16.1641 | -34.8532 | -25.5961
16.1702 | -34.8536 | -25.6038
16.1763 | -34.8541 | -25.6116
16.1824 | -34.8546 | -25.6193
16.1885 | -34.8551 | -25.627
16.1946 | -34.8556 | -25.6345
16.2007 | -34.8561 | -25.6423
16.2068 | -34.8566 | -25.6501
16.2129 | -34.8571 | -25.6579
16.219 | -34.8576 | -25.6654
16.2251 | -34.8581 | -25.6731
16.2312 | -34.8586 | -25.6809
16.2373 | -34.8591 | -25.6885
16.2434 | -34.8597 | -25.6964
16.2495 | -34.8602 | -25.7041
16.2556 | -34.8607 | -25.7119
16.2618 | -34.8612 | -25.7195
16.2679 | -34.8617 | -25.7274
16.274 | -34.8622 | -25.7353
16.2801 | -34.8627 | -25.7431
16.2862 | -34.8632 | -25.7509
16.2923 | -34.8638 | -25.7587
16.2984 | -34.8643 | -25.7664
16.3045 | -34.8648 | -25.7741
16.3106 | -34.8653 | -25.7819
16.3167 | -34.8658 | -25.7897
16.3228 | -34.8663 | -25.7973
16.3289 | -34.8668 | -25.8049
16.335 | -34.8673 | -25.8126
16.3411 | -34.8678 | -25.8203
16.3472 | -34.8683 | -25.8282
16.3533 | -34.8688 | -25.836
16.3594 | -34.8693 | -25.8434
16.3655 | -34.8698 | -25.8509
16.3716 | -34.8703 | -25.8585
16.3777 | -34.8708 | -25.8662
16.3838 | -34.8713 | -25.8739
16.3899 | -34.8718 | -25.8817
16.396 | -34.8723 | -25.8895
16.4021 | -34.8728 | -25.8971
16.4083 | -34.8733 | -25.9048
16.4144 | -34.8739 | -25.9127
16.4205 | -34.8744 | -25.9205
16.4266 | -34.8749 | -25.9281
16.4327 | -34.8754 | -25.9357
16.4388 | -34.8759 | -25.9434
16.4449 | -34.8764 | -25.9511
16.451 | -34.8769 | -25.9588
16.4571 | -34.8774 | -25.9665
16.4632 | -34.8779 | -25.974
16.4693 | -34.8784 | -25.9817
16.4754 | -34.8789 | -25.9893
16.4815 | -34.8794 | -25.9971
16.4876 | -34.8799 | -26.0047
16.4937 | -34.8804 | -26.0122
16.4998 | -34.8809 | -26.0198
16.5059 | -34.8815 | -26.0275
16.512 | -34.8819 | -26.0353
16.5181 | -34.8825 | -26.0431
16.5242 | -34.8829 | -26.051
16.5303 | -34.8834 | -26.0585
16.5364 | -34.8839 | -26.0661
16.5425 | -34.8844 | -26.0738
16.5487 | -34.8849 | -26.0813
16.5548 | -34.8854 | -26.089
16.5609 | -34.8859 | -26.0969
16.567 | -34.8864 | -26.1043
16.5731 | -34.887 | -26.1121
16.5792 | -34.8874 | -26.1198
16.5853 | -34.8879 | -26.1274
16.5914 | -34.8885 | -26.1353
16.5975 | -34.889 | -26.1432
16.6036 | -34.8895 | -26.1507
16.6097 | -34.89 | -26.1584
16.6158 | -34.8905 | -26.166
16.6219 | -34.891 | -26.1738
16.628 | -34.8915 | -26.1815
16.6341 | -34.892 | -26.1892
16.6402 | -34.8925 | -26.1967
16.6463 | -34.893 | -26.2044
16.6524 | -34.8935 | -26.2121
16.6585 | -34.894 | -26.2198
16.6646 | -34.8945 | -26.2273
16.6707 | -34.895 | -26.2349
16.6768 | -34.8955 | -26.2426
16.6829 | -34.896 | -26.2503
16.689 | -34.8966 | -26.2579
16.6952 | -34.8971 | -26.2657
16.7013 | -34.8976 | -26.2733
16.7074 | -34.8981 | -26.2809
16.7135 | -34.8986 | -26.2885
16.7196 | -34.8991 | -26.296
16.7257 | -34.8996 | -26.3036
16.7318 | -34.9001 | -26.3112
16.7379 | -34.9006 | -26.3187
16.744 | -34.9011 | -26.3262
16.7501 | -34.9016 | -26.3339
16.7562 | -34.9021 | -26.3415
16.7623 | -34.9025 | -26.349
16.7684 | -34.903 | -26.3567
16.7745 | -34.9035 | -26.3643
16.7806 | -34.9041 | -26.3715
16.7867 | -34.9046 | -26.3793
16.7928 | -34.9051 | -26.387
16.7989 | -34.9056 | -26.3945
16.805 | -34.9061 | -26.4021
16.8111 | -34.9066 | -26.4097
16.8172 | -34.907 | -26.4174
16.8233 | -34.9076 | -26.425
16.8294 | -34.908 | -26.4329
16.8356 | -34.9086 | -26.4405
16.8417 | -34.9091 | -26.4478
16.8478 | -34.9096 | -26.4554
16.8539 | -34.91 | -26.4629
16.86 | -34.9105 | -26.4707
16.8661 | -34.911 | -26.4782
16.8722 | -34.9115 | -26.4859
16.8783 | -34.912 | -26.4937
16.8844 | -34.9125 | -26.5013
16.8905 | -34.913 | -26.5089
16.8966 | -34.9136 | -26.5164
16.9027 | -34.9141 | -26.5237
16.9088 | -34.9146 | -26.5315
16.9149 | -34.9151 | -26.5391
16.921 | -34.9156 | -26.5468
16.9271 | -34.9161 | -26.5543
16.9332 | -34.9166 | -26.5622
16.9393 | -34.9171 | -26.57
16.9454 | -34.9175 | -26.5774
16.9515 | -34.918 | -26.5849
16.9576 | -34.9186 | -26.5925
16.9637 | -34.9191 | -26.6
16.9698 | -34.9196 | -26.6076
16.9759 | -34.9201 | -26.6152
16.9821 | -34.9206 | -26.6229
16.9882 | -34.9211 | -26.6304
16.9943 | -34.9216 | -26.6381
17.0004 | -34.9221 | -26.6458
17.0065 | -34.9226 | -26.6533
17.0126 | -34.9231 | -26.6608
17.0187 | -34.9236 | -26.6686
17.0248 | -34.9242 | -26.6762
17.0309 | -34.9247 | -26.6838
17.037 | -34.9252 | -26.6913
17.0431 | -34.9256 | -26.6988
17.0492 | -34.9261 | -26.7064
17.0553 | -34.9266 | -26.7142
17.0614 | -34.9271 | -26.7217
17.0675 | -34.9276 | -26.7292
17.0736 | -34.9281 | -26.7369
17.0797 | -34.9286 | -26.7445
17.0858 | -34.9291 | -26.7522
17.0919 | -34.9296 | -26.7598
17.098 | -34.9301 | -26.7671
17.1041 | -34.9306 | -26.7747
17.1102 | -34.9311 | -26.7821
17.1163 | -34.9316 | -26.7897
17.1225 | -34.9322 | -26.7971
17.1286 | -34.9327 | -26.8045
17.1347 | -34.9332 | -26.8122
17.1408 | -34.9337 | -26.8199
17.1469 | -34.9342 | -26.8273
17.153 | -34.9347 | -26.8349
17.1591 | -34.9352 | -26.8423
17.1652 | -34.9357 | -26.8497
17.1713 | -34.9362 | -26.8572
17.1774 | -34.9368 | -26.8647
17.1835 | -34.9372 | -26.8722
17.1896 | -34.9377 | -26.8797
17.1957 | -34.9383 | -26.8872
17.2018 | -34.9388 | -26.8946
17.2079 | -34.9393 | -26.9022
17.214 | -34.9398 | -26.9097
17.2201 | -34.9403 | -26.9174
17.2262 | -34.9408 | -26.9251
17.2323 | -34.9413 | -26.9329
17.2384 | -34.9418 | -26.9404
17.2445 | -34.9423 | -26.9477
17.2506 | -34.9428 | -26.9552
17.2567 | -34.9433 | -26.9626
17.2628 | -34.9438 | -26.9703
17.269 | -34.9443 | -26.9778
17.2751 | -34.9448 | -26.9854
17.2812 | -34.9452 | -26.993
17.2873 | -34.9457 | -27.0005
17.2934 | -34.9462 | -27.0081
17.2995 | -34.9467 | -27.0156
17.3056 | -34.9472 | -27.0232
17.3117 | -34.9477 | -27.0309
17.3178 | -34.9482 | -27.0385
17.3239 | -34.9487 | -27.0461
17.33 | -34.9492 | -27.0538
17.3361 | -34.9497 | -27.0613
17.3422 | -34.9501 | -27.0689
17.3483 | -34.9506 | -27.0765
17.3544 | -34.9511 | -27.0841
17.3605 | -34.9516 | -27.0917
17.3666 | -34.9521 | -27.0992
17.3727 | -34.9526 | -27.1069
17.3788 | -34.9531 | -27.1143
17.3849 | -34.9536 | -27.1218
17.391 | -34.954 | -27.1295
17.3971 | -34.9545 | -27.1368
17.4032 | -34.955 | -27.1442
17.4094 | -34.9555 | -27.1516
17.4155 | -34.956 | -27.1592
17.4216 | -34.9565 | -27.1666
17.4277 | -34.957 | -27.1743
17.4338 | -34.9575 | -27.1819
17.4399 | -34.958 | -27.1896
17.446 | -34.9584 | -27.1968
17.4521 | -34.9589 | -27.2043
17.4582 | -34.9594 | -27.2118
17.4643 | -34.9599 | -27.2192
17.4704 | -34.9604 | -27.2267
17.4765 | -34.9609 | -27.2343
17.4826 | -34.9614 | -27.2417
17.4887 | -34.9619 | -27.2494
17.4948 | -34.9623 | -27.2569
17.5009 | -34.9628 | -27.2643
17.507 | -34.9633 | -27.2719
17.5131 | -34.9638 | -27.2795
17.5192 | -34.9642 | -27.2869
17.5253 | -34.9647 | -27.2943
17.5314 | -34.9652 | -27.3019
17.5375 | -34.9657 | -27.3093
17.5436 | -34.9662 | -27.317
17.5497 | -34.9667 | -27.3245
17.5559 | -34.9672 | -27.3319
17.562 | -34.9677 | -27.3393
17.5681 | -34.9682 | -27.3466
17.5742 | -34.9687 | -27.3542
17.5803 | -34.9692 | -27.3616
17.5864 | -34.9697 | -27.369
17.5925 | -34.9702 | -27.3765
17.5986 | -34.9707 | -27.3839
17.6047 | -34.9712 | -27.3915
17.6108 | -34.9716 | -27.3989
17.6169 | -34.9721 | -27.4063
17.623 | -34.9726 | -27.4137
17.6291 | -34.9731 | -27.421
17.6352 | -34.9736 | -27.4285
17.6413 | -34.9741 | -27.4361
17.6474 | -34.9746 | -27.4436
17.6535 | -34.9751 | -27.451
17.6596 | -34.9756 | -27.4585
17.6657 | -34.9761 | -27.4661
17.6718 | -34.9766 | -27.4736
17.6779 | -34.9771 | -27.481
17.684 | -34.9776 | -27.4885
17.6901 | -34.9781 | -27.4959
17.6963 | -34.9785 | -27.503
17.7024 | -34.979 | -27.5104
17.7085 | -34.9795 | -27.5176
17.7146 | -34.98 | -27.525
17.7207 | -34.9805 | -27.5323
17.7268 | -34.981 | -27.5397
17.7329 | -34.9815 | -27.5473
17.739 | -34.982 | -27.5548
17.7451 | -34.9825 | -27.5621
17.7512 | -34.983 | -27.5696
17.7573 | -34.9835 | -27.5772
17.7634 | -34.9839 | -27.5843
17.7695 | -34.9844 | -27.5917
17.7756 | -34.9849 | -27.5991
17.7817 | -34.9854 | -27.6065
17.7878 | -34.9859 | -27.6141
17.7939 | -34.9864 | -27.6216
17.8 | -34.9869 | -27.6288
17.8061 | -34.9874 | -27.6363
17.8122 | -34.9879 | -27.6437
17.8183 | -34.9884 | -27.6512
17.8244 | -34.9888 | -27.6586
17.8305 | -34.9893 | -27.6661
17.8366 | -34.9898 | -27.6733
17.8428 | -34.9903 | -27.6807
17.8489 | -34.9908 | -27.688
17.855 | -34.9913 | -27.6953
17.8611 | -34.9918 | -27.7028
17.8672 | -34.9923 | -27.7101
17.8733 | -34.9928 | -27.7175
17.8794 | -34.9933 | -27.725
17.8855 | -34.9937 | -27.7325
17.8916 | -34.9942 | -27.74
17.8977 | -34.9947 | -27.7474
17.9038 | -34.9952 | -27.7547
17.9099 | -34.9956 | -27.7622
17.916 | -34.9961 | -27.7697
17.9221 | -34.9966 | -27.7771
17.9282 | -34.9971 | -27.7845
17.9343 | -34.9976 | -27.7919
17.9404 | -34.9981 | -27.799
17.9465 | -34.9986 | -27.8064
17.9526 | -34.999 | -27.8139
17.9587 | -34.9995 | -27.8213
17.9648 | -35 | -27.8288
17.9709 | -35.0005 | -27.8362
17.977 | -35.001 | -27.8436
17.9832 | -35.0014 | -27.851
17.9893 | -35.0019 | -27.8584
17.9954 | -35.0024 | -27.8658
18.0015 | -35.0029 | -27.8732
18.0076 | -35.0033 | -27.8804
18.0137 | -35.0038 | -27.8877
18.0198 | -35.0043 | -27.8951
18.0259 | -35.0047 | -27.9027
18.032 | -35.0052 | -27.91
18.0381 | -35.0057 | -27.9175
18.0442 | -35.0062 | -27.9247
18.0503 | -35.0066 | -27.9321
18.0564 | -35.0071 | -27.9395
18.0625 | -35.0076 | -27.947
18.0686 | -35.0081 | -27.9544
18.0747 | -35.0086 | -27.9618
18.0808 | -35.009 | -27.9689
18.0869 | -35.0095 | -27.9765
18.093 | -35.01 | -27.9839
18.0991 | -35.0105 | -27.9915
18.1052 | -35.0109 | -27.999
18.1113 | -35.0114 | -28.0064
18.1174 | -35.0119 | -28.0137
18.1236 | -35.0124 | -28.0212
18.1297 | -35.0129 | -28.0285
18.1358 | -35.0133 | -28.0359
18.1419 | -35.0138 | -28.0431
18.148 | -35.0143 | -28.0504
18.1541 | -35.0148 | -28.0576
18.1602 | -35.0153 | -28.0649
18.1663 | -35.0157 | -28.0721
18.1724 | -35.0162 | -28.0796
18.1785 | -35.0167 | -28.0867
18.1846 | -35.0171 | -28.094
18.1907 | -35.0176 | -28.1014
18.1968 | -35.0181 | -28.1086
18.2029 | -35.0186 | -28.1162
18.209 | -35.019 | -28.1237
18.2151 | -35.0195 | -28.1312
18.2212 | -35.02 | -28.1385
18.2273 | -35.0205 | -28.1459
18.2334 | -35.021 | -28.1534
18.2395 | -35.0214 | -28.1607
18.2456 | -35.0219 | -28.168
18.2517 | -35.0224 | -28.1753
18.2578 | -35.0228 | -28.1825
18.2639 | -35.0233 | -28.1899
18.2701 | -35.0238 | -28.1971
18.2762 | -35.0243 | -28.2043
18.2823 | -35.0248 | -28.2114
18.2884 | -35.0253 | -28.2186
18.2945 | -35.0257 | -28.2255
18.3006 | -35.0262 | -28.233
18.3067 | -35.0267 | -28.2402
18.3128 | -35.0271 | -28.2476
18.3189 | -35.0276 | -28.2548
18.325 | -35.028 | -28.2618
18.3311 | -35.0285 | -28.2691
18.3372 | -35.0289 | -28.2765
18.3433 | -35.0294 | -28.2838
18.3494 | -35.0299 | -28.2911
18.3555 | -35.0303 | -28.2983
18.3616 | -35.0308 | -28.3052
18.3677 | -35.0313 | -28.3125
18.3738 | -35.0318 | -28.3197
18.3799 | -35.0322 | -28.3269
18.386 | -35.0327 | -28.3343
18.3921 | -35.0331 | -28.3416
18.3982 | -35.0336 | -28.3486
18.4043 | -35.034 | -28.3559
18.4105 | -35.0345 | -28.3633
18.4166 | -35.035 | -28.3702
18.4227 | -35.0354 | -28.3775
18.4288 | -35.0359 | -28.3847
18.4349 | -35.0363 | -28.3917
18.441 | -35.0368 | -28.3989
18.4471 | -35.0373 | -28.4059
18.4532 | -35.0377 | -28.4132
18.4593 | -35.0382 | -28.4205
18.4654 | -35.0387 | -28.4275
18.4715 | -35.0392 | -28.4351
18.4776 | -35.0396 | -28.4425
18.4837 | -35.0401 | -28.45
18.4898 | -35.0405 | -28.4574
18.4959 | -35.041 | -28.465
18.502 | -35.0415 | -28.4721
18.5081 | -35.0419 | -28.4794
18.5142 | -35.0424 | -28.4868
18.5203 | -35.0429 | -28.4941
18.5264 | -35.0433 | -28.5012
18.5325 | -35.0438 | -28.5084
18.5386 | -35.0443 | -28.5157
18.5447 | -35.0448 | -28.5231
18.5508 | -35.0452 | -28.5304
18.557 | -35.0457 | -28.5378
18.5631 | -35.0462 | -28.5451
18.5692 | -35.0466 | -28.5525
18.5753 | -35.0471 | -28.5595
18.5814 | -35.0475 | -28.5667
18.5875 | -35.048 | -28.5739
18.5936 | -35.0484 | -28.5811
18.5997 | -35.0489 | -28.5886
18.6058 | -35.0493 | -28.5957
18.6119 | -35.0498 | -28.6027
18.618 | -35.0503 | -28.61
18.6241 | -35.0507 | -28.6173
18.6302 | -35.0512 | -28.6247
18.6363 | -35.0517 | -28.6319
18.6424 | -35.0522 | -28.6393
18.6485 | -35.0526 | -28.6464
18.6546 | -35.0531 | -28.6537
18.6607 | -35.0535 | -28.6609
18.6668 | -35.054 | -28.6681
18.6729 | -35.0545 | -28.6753
18.679 | -35.0549 | -28.6822
18.6851 | -35.0553 | -28.6893
18.6912 | -35.0558 | -28.6963
18.6974 | -35.0563 | -28.7033
18.7035 | -35.0567 | -28.7105
18.7096 | -35.0572 | -28.7177
18.7157 | -35.0577 | -28.7247
18.7218 | -35.0581 | -28.7318
18.7279 | -35.0586 | -28.739
18.734 | -35.0591 | -28.7465
18.7401 | -35.0596 | -28.7536
18.7462 | -35.06 | -28.7605
18.7523 | -35.0605 | -28.7679
18.7584 | -35.061 | -28.7751
18.7645 | -35.0614 | -28.7823
18.7706 | -35.0619 | -28.7894
18.7767 | -35.0623 | -28.7966
18.7828 | -35.0628 | -28.8036
18.7889 | -35.0632 | -28.8108
18.795 | -35.0637 | -28.8181
18.8011 | -35.0641 | -28.8253
18.8072 | -35.0646 | -28.8323
18.8133 | -35.065 | -28.8394
18.8194 | -35.0655 | -28.8463
18.8255 | -35.0659 | -28.8534
18.8316 | -35.0663 | -28.8607
18.8377 | -35.0668 | -28.8678
18.8439 | -35.0672 | -28.8747
18.85 | -35.0677 | -28.8819
18.8561 | -35.0681 | -28.8891
18.8622 | -35.0686 | -28.8963
18.8683 | -35.069 | -28.9034
18.8744 | -35.0695 | -28.9106
18.8805 | -35.07 | -28.918
18.8866 | -35.0704 | -28.9253
18.8927 | -35.0709 | -28.9321
18.8988 | -35.0713 | -28.939
18.9049 | -35.0718 | -28.946
18.911 | -35.0722 | -28.9532
18.9171 | -35.0726 | -28.9604
18.9232 | -35.0731 | -28.967
18.9293 | -35.0736 | -28.9742
18.9354 | -35.074 | -28.9814
18.9415 | -35.0745 | -28.9883
18.9476 | -35.0749 | -28.9952
18.9537 | -35.0754 | -29.0024
18.9598 | -35.0758 | -29.0091
18.9659 | -35.0763 | -29.0164
18.972 | -35.0767 | -29.0237
18.9781 | -35.0772 | -29.0306
18.9843 | -35.0776 | -29.0377
18.9904 | -35.0781 | -29.0449
18.9965 | -35.0786 | -29.0519
19.0026 | -35.079 | -29.059
19.0087 | -35.0795 | -29.0661
19.0148 | -35.08 | -29.073
19.0209 | -35.0804 | -29.0802
19.027 | -35.0809 | -29.087
19.0331 | -35.0814 | -29.094
19.0392 | -35.0818 | -29.1009
19.0453 | -35.0822 | -29.108
19.0514 | -35.0827 | -29.1149
19.0575 | -35.0831 | -29.1224
19.0636 | -35.0835 | -29.1292
19.0697 | -35.084 | -29.1362
19.0758 | -35.0844 | -29.143
19.0819 | -35.0849 | -29.15
19.088 | -35.0854 | -29.157
19.0941 | -35.0858 | -29.1642
19.1002 | -35.0863 | -29.1707
19.1063 | -35.0867 | -29.1779
19.1124 | -35.0872 | -29.185
19.1185 | -35.0876 | -29.1922
19.1246 | -35.0881 | -29.1991
19.1308 | -35.0885 | -29.2062
19.1369 | -35.089 | -29.2129
19.143 | -35.0894 | -29.22
19.1491 | -35.0899 | -29.2271
19.1552 | -35.0903 | -29.2342
19.1613 | -35.0907 | -29.2414
19.1674 | -35.0912 | -29.2484
19.1735 | -35.0916 | -29.255
19.1796 | -35.0921 | -29.2621
19.1857 | -35.0925 | -29.2693
19.1918 | -35.093 | -29.2764
19.1979 | -35.0934 | -29.2835
19.204 | -35.0939 | -29.2904
19.2101 | -35.0943 | -29.2973
19.2162 | -35.0948 | -29.3041
19.2223 | -35.0952 | -29.3111
19.2284 | -35.0956 | -29.3183
19.2345 | -35.096 | -29.3254
19.2406 | -35.0964 | -29.3318
19.2467 | -35.0969 | -29.3389
19.2528 | -35.0974 | -29.3462
19.2589 | -35.0978 | -29.353
19.265 | -35.0982 | -29.3601
19.2712 | -35.0987 | -29.367
19.2773 | -35.0991 | -29.3738
19.2834 | -35.0995 | -29.3807
19.2895 | -35.0999 | -29.3877
19.2956 | -35.1003 | -29.3948
19.3017 | -35.1007 | -29.4016
19.3078 | -35.1011 | -29.4085
19.3139 | -35.1016 | -29.4155
19.32 | -35.1021 | -29.4226
19.3261 | -35.1025 | -29.4294
19.3322 | -35.103 | -29.4362
19.3383 | -35.1034 | -29.4432
19.3444 | -35.1038 | -29.4504
19.3505 | -35.1043 | -29.4572
19.3566 | -35.1047 | -29.4639
19.3627 | -35.1051 | -29.4709
19.3688 | -35.1056 | -29.4779
19.3749 | -35.106 | -29.4848
19.381 | -35.1065 | -29.4912
19.3871 | -35.1069 | -29.4981
19.3932 | -35.1073 | -29.5056
19.3993 | -35.1077 | -29.5126
19.4054 | -35.1082 | -29.5195
19.4115 | -35.1086 | -29.5263
19.4177 | -35.1091 | -29.5332
19.4238 | -35.1095 | -29.54
19.4299 | -35.1099 | -29.5469
19.436 | -35.1103 | -29.5539
19.4421 | -35.1108 | -29.5606
19.4482 | -35.1112 | -29.5673
19.4543 | -35.1117 | -29.5737
19.4604 | -35.1121 | -29.5807
19.4665 | -35.1125 | -29.5877
19.4726 | -35.1129 | -29.5948
19.4787 | -35.1133 | -29.6016
19.4848 | -35.1137 | -29.6085
19.4909 | -35.1141 | -29.6152
19.497 | -35.1145 | -29.6223
19.5031 | -35.115 | -29.629
19.5092 | -35.1154 | -29.6357
19.5153 | -35.1158 | -29.6425
19.5214 | -35.1163 | -29.6493
19.5275 | -35.1167 | -29.656
19.5336 | -35.1171 | -29.6625
19.5397 | -35.1175 | -29.6694
19.5458 | -35.118 | -29.6762
19.5519 | -35.1184 | -29.6831
19.5581 | -35.1188 | -29.6898
19.5642 | -35.1192 | -29.6966
19.5703 | -35.1196 | -29.7036
19.5764 | -35.1201 | -29.7107
19.5825 | -35.1205 | -29.7177
19.5886 | -35.121 | -29.7244
19.5947 | -35.1213 | -29.7307
19.6008 | -35.1218 | -29.7377
19.6069 | -35.1223 | -29.7443
19.613 | -35.1227 | -29.7512
19.6191 | -35.1231 | -29.7578
19.6252 | -35.1236 | -29.7646
19.6313 | -35.1241 | -29.7712
19.6374 | -35.1245 | -29.7779
19.6435 | -35.1249 | -29.7844
19.6496 | -35.1253 | -29.7913
19.6557 | -35.1257 | -29.7979
19.6618 | -35.1261 | -29.8046
19.6679 | -35.1265 | -29.8114
19.674 | -35.127 | -29.8184
19.6801 | -35.1274 | -29.8254
19.6862 | -35.1279 | -29.8321
19.6923 | -35.1283 | -29.8386
19.6984 | -35.1288 | -29.8451
19.7046 | -35.1293 | -29.8521
19.7107 | -35.1297 | -29.8591
19.7168 | -35.1302 | -29.8657
19.7229 | -35.1307 | -29.8725
19.729 | -35.1311 | -29.8793
19.7351 | -35.1314 | -29.8858
19.7412 | -35.1319 | -29.8927
19.7473 | -35.1323 | -29.8997
19.7534 | -35.1327 | -29.9065
19.7595 | -35.1331 | -29.9132
19.7656 | -35.1336 | -29.9199
19.7717 | -35.134 | -29.9263
19.7778 | -35.1345 | -29.9331
19.7839 | -35.1349 | -29.9396
19.79 | -35.1353 | -29.9462
19.7961 | -35.1358 | -29.9528
19.8022 | -35.1362 | -29.9593
19.8083 | -35.1367 | -29.9662
19.8144 | -35.1371 | -29.9731
19.8205 | -35.1377 | -29.9799
19.8266 | -35.1381 | -29.9863
19.8327 | -35.1385 | -29.993
19.8388 | -35.1389 | -29.9993
19.845 | -35.1393 | -30.0059
19.8511 | -35.1397 | -30.0132
19.8572 | -35.1401 | -30.0196
19.8633 | -35.1406 | -30.0264
19.8694 | -35.141 | -30.0331
19.8755 | -35.1415 | -30.0394
19.8816 | -35.1419 | -30.0467
19.8877 | -35.1423 | -30.0533
19.8938 | -35.1428 | -30.0602
19.8999 | -35.1432 | -30.0672
19.906 | -35.1436 | -30.0738
19.9121 | -35.1441 | -30.08
19.9182 | -35.1445 | -30.0868
19.9243 | -35.1449 | -30.0935
19.9304 | -35.1453 | -30.1001
19.9365 | -35.1458 | -30.1069
19.9426 | -35.1462 | -30.113
19.9487 | -35.1467 | -30.1198
19.9548 | -35.1471 | -30.1266
19.9609 | -35.1475 | -30.1332
19.967 | -35.1479 | -30.1396
19.9731 | -35.1483 | -30.1462
19.9792 | -35.1487 | -30.1528
19.9853 | -35.1491 | -30.1596
19.9915 | -35.1495 | -30.1664
19.9976 | -35.15 | -30.173
20.0037 | -35.1505 | -30.1797
20.0098 | -35.1509 | -30.1863
20.0159 | -35.1513 | -30.1927
20.022 | -35.1517 | -30.1998
20.0281 | -35.1521 | -30.2066
20.0342 | -35.1525 | -30.2134
20.0403 | -35.153 | -30.2201
20.0464 | -35.1534 | -30.2268
20.0525 | -35.1539 | -30.2337
20.0586 | -35.1543 | -30.2403
20.0647 | -35.1547 | -30.2471
20.0708 | -35.1552 | -30.2541
20.0769 | -35.1556 | -30.2608
20.083 | -35.156 | -30.2673
20.0891 | -35.1565 | -30.2743
20.0952 | -35.1568 | -30.2809
20.1013 | -35.1573 | -30.2875
20.1074 | -35.1578 | -30.2944
20.1135 | -35.1582 | -30.3013
20.1196 | -35.1587 | -30.3079
20.1257 | -35.1592 | -30.3144
20.1319 | -35.1596 | -30.3206
20.138 | -35.16 | -30.327
20.1441 | -35.1605 | -30.3338
20.1502 | -35.161 | -30.3406
20.1563 | -35.1614 | -30.3462
20.1624 | -35.1619 | -30.3524
20.1685 | -35.1623 | -30.3587
20.1746 | -35.1627 | -30.3652
20.1807 | -35.1631 | -30.3718
20.1868 | -35.1635 | -30.3785
20.1929 | -35.1639 | -30.3845
20.199 | -35.1643 | -30.391
20.2051 | -35.1647 | -30.3975
20.2112 | -35.1651 | -30.4044
20.2173 | -35.1655 | -30.4113
20.2234 | -35.1659 | -30.4173
20.2295 | -35.1663 | -30.4232
20.2356 | -35.1668 | -30.4301
20.2417 | -35.1672 | -30.437
20.2478 | -35.1676 | -30.4436
20.2539 | -35.168 | -30.4502
20.26 | -35.1686 | -30.4562
20.2661 | -35.169 | -30.463
20.2723 | -35.1695 | -30.4694
20.2784 | -35.1699 | -30.476
20.2845 | -35.1704 | -30.4823
20.2906 | -35.1707 | -30.4886
20.2967 | -35.1711 | -30.4944
20.3028 | -35.1715 | -30.5012
20.3089 | -35.1719 | -30.508
20.315 | -35.1724 | -30.5141
20.3211 | -35.1728 | -30.5203
20.3272 | -35.1732 | -30.5265
20.3333 | -35.1737 | -30.5321
20.3394 | -35.1741 | -30.5392
20.3455 | -35.1746 | -30.5457
20.3516 | -35.1749 | -30.5519
20.3577 | -35.1753 | -30.5584
20.3638 | -35.1758 | -30.565
20.3699 | -35.1763 | -30.5717
20.376 | -35.1767 | -30.5781
20.3821 | -35.1772 | -30.5846
20.3882 | -35.1776 | -30.5912
20.3943 | -35.1781 | -30.5981
20.4004 | -35.1785 | -30.6036
20.4065 | -35.179 | -30.6102
20.4126 | -35.1794 | -30.6164
20.4188 | -35.1798 | -30.6233
20.4249 | -35.1802 | -30.63
20.431 | -35.1806 | -30.6368
20.4371 | -35.1811 | -30.643
20.4432 | -35.1816 | -30.6498
20.4493 | -35.182 | -30.6564
20.4554 | -35.1825 | -30.6628
20.4615 | -35.1829 | -30.6695
20.4676 | -35.1833 | -30.676
20.4737 | -35.1837 | -30.6818
20.4798 | -35.1841 | -30.6882
20.4859 | -35.1845 | -30.695
20.492 | -35.1849 | -30.7018
20.4981 | -35.1854 | -30.7086
20.5042 | -35.1858 | -30.7152
20.5103 | -35.1861 | -30.7214
20.5164 | -35.1865 | -30.7274
20.5225 | -35.1869 | -30.7342
20.5286 | -35.1873 | -30.7406
20.5347 | -35.1877 | -30.7473
20.5408 | -35.1881 | -30.7533
20.5469 | -35.1886 | -30.7598
20.553 | -35.189 | -30.7661
20.5592 | -35.1894 | -30.7729
20.5653 | -35.1898 | -30.7795
20.5714 | -35.1902 | -30.7864
20.5775 | -35.1906 | -30.7927
20.5836 | -35.1911 | -30.7992
20.5897 | -35.1914 | -30.8058
20.5958 | -35.1918 | -30.8131
20.6019 | -35.1922 | -30.8197
20.608 | -35.1927 | -30.8261
20.6141 | -35.1931 | -30.8321
20.6202 | -35.1935 | -30.8384
20.6263 | -35.194 | -30.8444
20.6324 | -35.1945 | -30.8507
20.6385 | -35.1949 | -30.857
20.6446 | -35.1954 | -30.8634
20.6507 | -35.1958 | -30.8698
20.6568 | -35.1963 | -30.8762
20.6629 | -35.1967 | -30.8828
20.669 | -35.1971 | -30.8896
20.6751 | -35.1976 | -30.8957
20.6812 | -35.1979 | -30.9024
20.6873 | -35.1983 | -30.9095
20.6934 | -35.1988 | -30.9156
20.6995 | -35.1992 | -30.9221
20.7057 | -35.1996 | -30.9289
20.7118 | -35.2001 | -30.9352
20.7179 | -35.2004 | -30.9419
20.724 | -35.2008 | -30.9484
20.7301 | -35.2012 | -30.9552
20.7362 | -35.2016 | -30.9614
20.7423 | -35.202 | -30.9681
20.7484 | -35.2025 | -30.9746
20.7545 | -35.2031 | -30.9816
20.7606 | -35.2034 | -30.9889
20.7667 | -35.2039 | -30.9961
20.7728 | -35.2042 | -31.0028
20.7789 | -35.2046 | -31.0092
20.785 | -35.205 | -31.016
20.7911 | -35.2054 | -31.0228
20.7972 | -35.2058 | -31.0294
20.8033 | -35.2062 | -31.036
20.8094 | -35.2067 | -31.042
20.8155 | -35.2071 | -31.0484
20.8216 | -35.2076 | -31.0549
20.8277 | -35.2081 | -31.0608
20.8338 | -35.2086 | -31.0675
20.8399 | -35.209 | -31.0739
20.8461 | -35.2093 | -31.0797
20.8522 | -35.2098 | -31.0859
20.8583 | -35.2102 | -31.0919
20.8644 | -35.2106 | -31.0982
20.8705 | -35.211 | -31.1046
20.8766 | -35.2115 | -31.1109
20.8827 | -35.2119 | -31.1174
20.8888 | -35.2123 | -31.1245
20.8949 | -35.2128 | -31.1307
20.901 | -35.2131 | -31.1371
20.9071 | -35.2135 | -31.1438
20.9132 | -35.2139 | -31.1501
20.9193 | -35.2143 | -31.1566
20.9254 | -35.2148 | -31.1635
20.9315 | -35.2153 | -31.1693
20.9376 | -35.2157 | -31.175
20.9437 | -35.216 | -31.181
20.9498 | -35.2164 | -31.1873
20.9559 | -35.2169 | -31.1935
20.962 | -35.2173 | -31.1998
20.9681 | -35.2177 | -31.2048
20.9742 | -35.2181 | -31.211
20.9803 | -35.2186 | -31.2172
20.9864 | -35.219 | -31.2231
20.9926 | -35.2195 | -31.2292
20.9987 | -35.22 | -31.2356
21.0048 | -35.2204 | -31.2418
21.0109 | -35.2209 | -31.2475
21.017 | -35.2213 | -31.2537
21.0231 | -35.2218 | -31.2599
21.0292 | -35.2222 | -31.2663
21.0353 | -35.2227 | -31.2717
21.0414 | -35.2231 | -31.2778
21.0475 | -35.2234 | -31.2847
21.0536 | -35.2238 | -31.2913
21.0597 | -35.2243 | -31.2971
21.0658 | -35.2246 | -31.3026
21.0719 | -35.225 | -31.3083
21.078 | -35.2254 | -31.3149
21.0841 | -35.2259 | -31.3215
21.0902 | -35.2264 | -31.3281
21.0963 | -35.2269 | -31.3344
21.1024 | -35.2273 | -31.3404
21.1085 | -35.2278 | -31.3465
21.1146 | -35.2283 | -31.3531
21.1207 | -35.2286 | -31.3595
21.1268 | -35.229 | -31.3658
21.133 | -35.2295 | -31.3721
21.1391 | -35.2299 | -31.3783
21.1452 | -35.2304 | -31.3845
21.1513 | -35.2308 | -31.3906
21.1574 | -35.2313 | -31.3965
21.1635 | -35.2318 | -31.403
21.1696 | -35.2322 | -31.4089
21.1757 | -35.2326 | -31.415
21.1818 | -35.233 | -31.4214
21.1879 | -35.2335 | -31.4272
21.194 | -35.2339 | -31.4334
21.2001 | -35.2343 | -31.44
21.2062 | -35.2347 | -31.4462
21.2123 | -35.2352 | -31.4521
21.2184 | -35.2357 | -31.4584
21.2245 | -35.2362 | -31.4645
21.2306 | -35.2366 | -31.4706
21.2367 | -35.2371 | -31.477
21.2428 | -35.2375 | -31.4829
21.2489 | -35.2379 | -31.4891
21.255 | -35.2383 | -31.4958
21.2611 | -35.2388 | -31.5021
21.2672 | -35.2393 | -31.5083
21.2733 | -35.2397 | -31.5141
21.2795 | -35.2401 | -31.5204
21.2856 | -35.2406 | -31.5268
21.2917 | -35.241 | -31.5336
21.2978 | -35.2415 | -31.54
21.3039 | -35.2419 | -31.5456
21.31 | -35.2423 | -31.5524
21.3161 | -35.2428 | -31.5588
21.3222 | -35.2433 | -31.5649
21.3283 | -35.2439 | -31.5708
21.3344 | -35.2443 | -31.5771
21.3405 | -35.2447 | -31.5831
21.3466 | -35.2451 | -31.589
21.3527 | -35.2455 | -31.5957
21.3588 | -35.2459 | -31.6017
21.3649 | -35.2463 | -31.6072
21.371 | -35.2468 | -31.6135
21.3771 | -35.2472 | -31.6196
21.3832 | -35.2477 | -31.6259
21.3893 | -35.2484 | -31.6327
21.3954 | -35.2488 | -31.6386
21.4015 | -35.2493 | -31.6449
21.4076 | -35.2498 | -31.6508
21.4137 | -35.2502 | -31.6571
21.4199 | -35.2506 | -31.6631
21.426 | -35.2511 | -31.6689
21.4321 | -35.2515 | -31.6752
21.4382 | -35.252 | -31.6814
21.4443 | -35.2525 | -31.688
21.4504 | -35.2529 | -31.6939
21.4565 | -35.2534 | -31.6996
21.4626 | -35.2539 | -31.7059
21.4687 | -35.2544 | -31.7119
21.4748 | -35.2549 | -31.7178
21.4809 | -35.2554 | -31.7238
21.487 | -35.2558 | -31.7301
21.4931 | -35.2563 | -31.7355
21.4992 | -35.2568 | -31.7413
21.5053 | -35.2572 | -31.7476
21.5114 | -35.2577 | -31.7539
21.5175 | -35.2581 | -31.7601
21.5236 | -35.2585 | -31.7665
21.5297 | -35.259 | -31.7721
21.5358 | -35.2593 | -31.779
21.5419 | -35.2597 | -31.7855
21.548 | -35.2601 | -31.7923
21.5541 | -35.2605 | -31.7979
21.5602 | -35.2611 | -31.8041
21.5664 | -35.2616 | -31.8091
21.5725 | -35.2621 | -31.8159
21.5786 | -35.2626 | -31.8225
21.5847 | -35.2631 | -31.8288
21.5908 | -35.2636 | -31.8342
21.5969 | -35.2643 | -31.8402
21.603 | -35.2648 | -31.8463
21.6091 | -35.2652 | -31.853
21.6152 | -35.2655 | -31.8594
21.6213 | -35.2659 | -31.8654
21.6274 | -35.2663 | -31.871
21.6335 | -35.2667 | -31.8757
21.6396 | -35.2672 | -31.8815
21.6457 | -35.2676 | -31.8867
21.6518 | -35.268 | -31.8933
21.6579 | -35.2684 | -31.8988
21.664 | -35.2688 | -31.9047
21.6701 | -35.2693 | -31.912
21.6762 | -35.2698 | -31.9183
21.6823 | -35.2703 | -31.9245
21.6884 | -35.2707 | -31.9301
21.6945 | -35.2712 | -31.9367
21.7006 | -35.2717 | -31.9429
21.7068 | -35.2725 | -31.9483
21.7129 | -35.2729 | -31.9542
21.719 | -35.2735 | -31.9607
21.7251 | -35.274 | -31.9659
21.7312 | -35.2745 | -31.9717
21.7373 | -35.275 | -31.977
21.7434 | -35.2754 | -31.9833
21.7495 | -35.2759 | -31.9903
21.7556 | -35.2764 | -31.9968
21.7617 | -35.2768 | -32.0028
21.7678 | -35.2772 | -32.0085
21.7739 | -35.2778 | -32.0154
21.78 | -35.2783 | -32.0215
21.7861 | -35.2787 | -32.0277
21.7922 | -35.2791 | -32.0334
21.7983 | -35.2795 | -32.0399
21.8044 | -35.28 | -32.0459
21.8105 | -35.2803 | -32.0528
21.8166 | -35.2807 | -32.0588
21.8227 | -35.2812 | -32.0657
21.8288 | -35.2816 | -32.0714
21.8349 | -35.282 | -32.0763
21.841 | -35.2823 | -32.0828
21.8471 | -35.2828 | -32.0901
21.8533 | -35.2833 | -32.0959
21.8594 | -35.2838 | -32.1012
21.8655 | -35.2841 | -32.1083
21.8716 | -35.2846 | -32.1142
21.8777 | -35.2848 | -32.1208
21.8838 | -35.2853 | -32.1255
21.8899 | -35.2858 | -32.1308
21.896 | -35.2862 | -32.1379
21.9021 | -35.2867 | -32.1435
21.9082 | -35.2871 | -32.1498
21.9143 | -35.2877 | -32.1552
21.9204 | -35.2881 | -32.1614
21.9265 | -35.2885 | -32.167
21.9326 | -35.2889 | -32.1736
21.9387 | -35.2892 | -32.1802
21.9448 | -35.2897 | -32.1852
21.9509 | -35.2902 | -32.1892
21.957 | -35.2906 | -32.1953
21.9631 | -35.2909 | -32.2012
21.9692 | -35.2913 | -32.2064
21.9753 | -35.2917 | -32.2114
21.9814 | -35.2923 | -32.2179
21.9875 | -35.2931 | -32.2249
21.9937 | -35.2936 | -32.232
21.9998 | -35.2939 | -32.2387
22.0059 | -35.2943 | -32.2435
22.012 | -35.2947 | -32.2496
22.0181 | -35.2953 | -32.2562
22.0242 | -35.2958 | -32.2634
22.0303 | -35.2963 | -32.2686
22.0364 | -35.2967 | -32.2743
22.0425 | -35.2974 | -32.2796
22.0486 | -35.2979 | -32.2847
22.0547 | -35.2984 | -32.2903
22.0608 | -35.2987 | -32.2971
22.0669 | -35.2992 | -32.3019
22.073 | -35.2996 | -32.3081
22.0791 | -35.3001 | -32.3126
22.0852 | -35.3006 | -32.3205
22.0913 | -35.3009 | -32.3253
22.0974 | -35.3013 | -32.332
22.1035 | -35.3018 | -32.3381
22.1096 | -35.3023 | -32.3433
22.1157 | -35.3027 | -32.3494
22.1218 | -35.303 | -32.3564
22.1279 | -35.3033 | -32.3617
22.134 | -35.3037 | -32.3691
22.1402 | -35.3041 | -32.3761
22.1463 | -35.3044 | -32.3828
22.1524 | -35.3051 | -32.3885
22.1585 | -35.3054 | -32.3943
22.1646 | -35.306 | -32.3997
22.1707 | -35.3065 | -32.4058
22.1768 | -35.307 | -32.4121
22.1829 | -35.3073 | -32.4182
22.189 | -35.3079 | -32.4245
22.1951 | -35.3081 | -32.4302
22.2012 | -35.3086 | -32.4368
22.2073 | -35.3091 | -32.4428
22.2134 | -35.3095 | -32.4487
22.2195 | -35.3098 | -32.4547
22.2256 | -35.3103 | -32.4615
22.2317 | -35.3108 | -32.4663
22.2378 | -35.3113 | -32.4723
22.2439 | -35.3117 | -32.4781
22.25 | -35.3123 | -32.4841
22.2561 | -35.3126 | -32.49
22.2622 | -35.3131 | -32.4956
22.2683 | -35.3137 | -32.5008
22.2744 | -35.3141 | -32.5072
22.2806 | -35.3145 | -32.5137
22.2867 | -35.315 | -32.5197
22.2928 | -35.3156 | -32.5256
22.2989 | -35.3159 | -32.5323
22.305 | -35.3164 | -32.5389
22.3111 | -35.3169 | -32.546
22.3172 | -35.3173 | -32.5517
22.3233 | -35.3178 | -32.5573
22.3294 | -35.3182 | -32.5643
22.3355 | -35.3188 | -32.5711
22.3416 | -35.3194 | -32.5775
22.3477 | -35.3199 | -32.5834
22.3538 | -35.3203 | -32.5897
22.3599 | -35.3208 | -32.5958
22.366 | -35.3211 | -32.602
22.3721 | -35.3216 | -32.6063
22.3782 | -35.3219 | -32.6126
22.3843 | -35.3225 | -32.6189
22.3904 | -35.3231 | -32.6244
22.3965 | -35.3235 | -32.6293
22.4026 | -35.3241 | -32.6347
22.4087 | -35.3245 | -32.6407
22.4148 | -35.325 | -32.646
22.4209 | -35.3254 | -32.6507
22.4271 | -35.3258 | -32.6576
22.4332 | -35.3262 | -32.6627
22.4393 | -35.3266 | -32.6695
22.4454 | -35.3272 | -32.6774
22.4515 | -35.3275 | -32.6842
22.4576 | -35.3281 | -32.6911
22.4637 | -35.3285 | -32.6966
22.4698 | -35.3291 | -32.702
22.4759 | -35.3293 | -32.707
22.482 | -35.3297 | -32.713
22.4881 | -35.3302 | -32.7198
22.4942 | -35.3307 | -32.7262
22.5003 | -35.3313 | -32.7329
22.5064 | -35.3317 | -32.7387
22.5125 | -35.332 | -32.745
22.5186 | -35.3324 | -32.7499
22.5247 | -35.3329 | -32.7544
22.5308 | -35.3335 | -32.7593
22.5369 | -35.334 | -32.7656
22.543 | -35.3346 | -32.7706
22.5491 | -35.3351 | -32.7766
22.5552 | -35.3356 | -32.7834
22.5613 | -35.3361 | -32.79
22.5675 | -35.3365 | -32.7973
22.5736 | -35.3369 | -32.8031
22.5797 | -35.3373 | -32.8098
22.5858 | -35.338 | -32.8163
22.5919 | -35.3384 | -32.8211
22.598 | -35.339 | -32.8271
22.6041 | -35.3395 | -32.8329
22.6102 | -35.3399 | -32.8383
22.6163 | -35.3404 | -32.843
22.6224 | -35.3408 | -32.8498
22.6285 | -35.3412 | -32.8552
22.6346 | -35.3416 | -32.8604
22.6407 | -35.342 | -32.8659
22.6468 | -35.3426 | -32.8706
22.6529 | -35.3429 | -32.8765
22.659 | -35.3432 | -32.8816
22.6651 | -35.3437 | -32.8883
22.6712 | -35.3441 | -32.8948
22.6773 | -35.3445 | -32.9014
22.6834 | -35.345 | -32.9073
22.6895 | -35.3456 | -32.9147
22.6956 | -35.346 | -32.9207
22.7017 | -35.3465 | -32.9274
22.7079 | -35.3469 | -32.9339
22.714 | -35.3473 | -32.94
22.7201 | -35.3478 | -32.9462
22.7262 | -35.3482 | -32.9515
22.7323 | -35.3488 | -32.9581
22.7384 | -35.3493 | -32.9637
22.7445 | -35.3498 | -32.9706
22.7506 | -35.3502 | -32.9758
22.7567 | -35.3506 | -32.9815
22.7628 | -35.3513 | -32.9862
22.7689 | -35.3518 | -32.992
22.775 | -35.3521 | -32.997
22.7811 | -35.3526 | -33.0036
22.7872 | -35.3532 | -33.0097
22.7933 | -35.3535 | -33.0159
22.7994 | -35.3541 | -33.0218
22.8055 | -35.3545 | -33.0281
22.8116 | -35.3551 | -33.0343
22.8177 | -35.3554 | -33.0407
22.8238 | -35.3559 | -33.0462
22.8299 | -35.3566 | -33.052
22.836 | -35.3571 | -33.0574
22.8421 | -35.3576 | -33.0633
22.8482 | -35.3579 | -33.0685
22.8544 | -35.3584 | -33.0742
22.8605 | -35.3589 | -33.0801
22.8666 | -35.3593 | -33.0857
22.8727 | -35.3597 | -33.0912
22.8788 | -35.3602 | -33.0968
22.8849 | -35.3607 | -33.1029
22.891 | -35.361 | -33.1084
22.8971 | -35.3615 | -33.114
22.9032 | -35.3619 | -33.1206
22.9093 | -35.3623 | -33.1266
22.9154 | -35.3627 | -33.1334
22.9215 | -35.3633 | -33.1402
22.9276 | -35.3637 | -33.1466
22.9337 | -35.3642 | -33.1519
22.9398 | -35.3647 | -33.1579
22.9459 | -35.3652 | -33.1639
22.952 | -35.3656 | -33.1689
22.9581 | -35.3659 | -33.1753
22.9642 | -35.3663 | -33.1805
22.9703 | -35.3663 | -33.1868
22.9764 | -35.3667 | -33.1932
22.9825 | -35.3672 | -33.1999
22.9886 | -35.3676 | -33.2063
22.9948 | -35.3681 | -33.212
23.0009 | -35.3685 | -33.2184
23.007 | -35.3689 | -33.2246
23.0131 | -35.3692 | -33.2308
23.0192 | -35.3695 | -33.2365
23.0253 | -35.3699 | -33.243
23.0314 | -35.3705 | -33.2482
23.0375 | -35.3709 | -33.2532
23.0436 | -35.3715 | -33.2588
23.0497 | -35.3721 | -33.2647
23.0558 | -35.3725 | -33.2704
23.0619 | -35.3729 | -33.2767
23.068 | -35.3733 | -33.2833
23.0741 | -35.3738 | -33.2883
23.0802 | -35.3742 | -33.2947
23.0863 | -35.3747 | -33.3009
23.0924 | -35.3751 | -33.3067
23.0985 | -35.3755 | -33.3122
23.1046 | -35.376 | -33.3181
23.1107 | -35.3763 | -33.3228
23.1168 | -35.3768 | -33.3288
23.1229 | -35.3773 | -33.3341
23.129 | -35.3779 | -33.34
23.1351 | -35.3784 | -33.3458
23.1413 | -35.3787 | -33.352
23.1474 | -35.379 | -33.3577
23.1535 | -35.3795 | -33.3634
23.1596 | -35.3798 | -33.3695
23.1657 | -35.3803 | -33.3742
23.1718 | -35.3809 | -33.3803
23.1779 | -35.3812 | -33.3853
23.184 | -35.3817 | -33.3915
23.1901 | -35.3821 | -33.3974
23.1962 | -35.3826 | -33.4033
23.2023 | -35.383 | -33.4088
23.2084 | -35.3835 | -33.4153
23.2145 | -35.3841 | -33.4207
23.2206 | -35.3845 | -33.4278
23.2267 | -35.3852 | -33.4334
23.2328 | -35.3857 | -33.4385
23.2389 | -35.3862 | -33.4436
23.245 | -35.3866 | -33.4498
23.2511 | -35.3872 | -33.4559
23.2572 | -35.3877 | -33.462
23.2633 | -35.3882 | -33.4677
23.2694 | -35.3885 | -33.4726
23.2755 | -35.3891 | -33.4785
23.2817 | -35.3896 | -33.4833
23.2878 | -35.39 | -33.4889
23.2939 | -35.3905 | -33.4945
23.3 | -35.3909 | -33.5009
23.3061 | -35.3914 | -33.5072
23.3122 | -35.3918 | -33.5129
23.3183 | -35.3923 | -33.5181
23.3244 | -35.3927 | -33.5246
23.3305 | -35.3932 | -33.5307
23.3366 | -35.3937 | -33.5367
23.3427 | -35.3941 | -33.5427
23.3488 | -35.3945 | -33.5482
23.3549 | -35.395 | -33.5543
23.361 | -35.3954 | -33.5607
23.3671 | -35.3959 | -33.5663
23.3732 | -35.3964 | -33.5721
23.3793 | -35.397 | -33.578
23.3854 | -35.3975 | -33.5837
23.3915 | -35.398 | -33.5907
23.3976 | -35.3984 | -33.5961
23.4037 | -35.3987 | -33.6015
23.4098 | -35.3992 | -33.6078
23.4159 | -35.3996 | -33.613
23.422 | -35.4001 | -33.6184
23.4282 | -35.4006 | -33.6237
23.4343 | -35.401 | -33.6294
23.4404 | -35.4014 | -33.6356
23.4465 | -35.4019 | -33.6413
23.4526 | -35.4024 | -33.6476
23.4587 | -35.4028 | -33.6531
23.4648 | -35.4033 | -33.6592
23.4709 | -35.4038 | -33.6655
23.477 | -35.4043 | -33.6717
23.4831 | -35.4047 | -33.6777
23.4892 | -35.4051 | -33.6835
23.4953 | -35.4055 | -33.6897
23.5014 | -35.4059 | -33.6952
23.5075 | -35.4064 | -33.7011
23.5136 | -35.4068 | -33.7073
23.5197 | -35.4073 | -33.7136
23.5258 | -35.4078 | -33.7197
23.5319 | -35.4083 | -33.7261
23.538 | -35.4088 | -33.7322
23.5441 | -35.4092 | -33.738
23.5502 | -35.4097 | -33.7441
23.5563 | -35.4102 | -33.7502
23.5624 | -35.4106 | -33.7566
23.5686 | -35.411 | -33.7625
23.5747 | -35.4115 | -33.7682
23.5808 | -35.4119 | -33.7736
23.5869 | -35.4123 | -33.7793
23.593 | -35.4128 | -33.785
23.5991 | -35.4133 | -33.7906
23.6052 | -35.4138 | -33.7962
23.6113 | -35.4143 | -33.8017
23.6174 | -35.4148 | -33.8077
23.6235 | -35.4152 | -33.8136
23.6296 | -35.4156 | -33.8195
23.6357 | -35.4161 | -33.8256
23.6418 | -35.4167 | -33.831
23.6479 | -35.417 | -33.836
23.654 | -35.4175 | -33.8429
23.6601 | -35.4178 | -33.8493
23.6662 | -35.4183 | -33.8547
23.6723 | -35.4188 | -33.8622
23.6784 | -35.4192 | -33.8681
23.6845 | -35.4197 | -33.8739
23.6906 | -35.4201 | -33.8795
23.6967 | -35.4206 | -33.8856
23.7028 | -35.4212 | -33.892
23.7089 | -35.4215 | -33.8978
23.7151 | -35.422 | -33.903
23.7212 | -35.4225 | -33.9089
23.7273 | -35.4229 | -33.9154
23.7334 | -35.4233 | -33.9212
23.7395 | -35.4238 | -33.9271
23.7456 | -35.4242 | -33.9327
23.7517 | -35.4247 | -33.9395
23.7578 | -35.425 | -33.9453
23.7639 | -35.4255 | -33.9514
23.77 | -35.4259 | -33.957
23.7761 | -35.4263 | -33.9629
23.7822 | -35.4268 | -33.9689
23.7883 | -35.4273 | -33.9749
23.7944 | -35.4276 | -33.981
23.8005 | -35.4281 | -33.987
23.8066 | -35.4285 | -33.9938
23.8127 | -35.4289 | -33.9988
23.8188 | -35.4294 | -34.005
23.8249 | -35.4298 | -34.0109
23.831 | -35.4303 | -34.017
23.8371 | -35.4307 | -34.0228
23.8432 | -35.4311 | -34.0288
23.8493 | -35.4316 | -34.0347
23.8555 | -35.432 | -34.041
23.8616 | -35.4325 | -34.0465
23.8677 | -35.433 | -34.0524
23.8738 | -35.4335 | -34.0587
23.8799 | -35.434 | -34.0647
23.886 | -35.4343 | -34.0706
23.8921 | -35.4347 | -34.0768
23.8982 | -35.435 | -34.0828
23.9043 | -35.4355 | -34.0886
23.9104 | -35.436 | -34.0948
23.9165 | -35.4364 | -34.1008
23.9226 | -35.4369 | -34.1068
23.9287 | -35.4373 | -34.1127
23.9348 | -35.4378 | -34.1186
23.9409 | -35.4382 | -34.124
23.947 | -35.4386 | -34.1299
23.9531 | -35.4392 | -34.1361
23.9592 | -35.4396 | -34.1418
23.9653 | -35.4401 | -34.148
23.9714 | -35.4405 | -34.1533
23.9775 | -35.4409 | -34.1587
23.9836 | -35.4414 | -34.1644
23.9897 | -35.4419 | -34.1707
23.9958 | -35.4424 | -34.177
24.002 | -35.4428 | -34.1826
24.0081 | -35.4432 | -34.1886
24.0142 | -35.4436 | -34.194
24.0203 | -35.444 | -34.1997
24.0264 | -35.4444 | -34.2054
24.0325 | -35.4449 | -34.2113
24.0386 | -35.4453 | -34.2174
24.0447 | -35.4458 | -34.2231
24.0508 | -35.4462 | -34.2292
24.0569 | -35.4467 | -34.2348
24.063 | -35.4471 | -34.2402
24.0691 | -35.4476 | -34.2465
24.0752 | -35.448 | -34.253
24.0813 | -35.4485 | -34.259
24.0874 | -35.4489 | -34.2649
24.0935 | -35.4493 | -34.2706
24.0996 | -35.4496 | -34.2768
24.1057 | -35.4502 | -34.2827
24.1118 | -35.4506 | -34.2882
24.1179 | -35.4511 | -34.2941
24.124 | -35.4515 | -34.2996
24.1301 | -35.4519 | -34.3055
24.1362 | -35.4523 | -34.3116
24.1424 | -35.4528 | -34.3174
24.1485 | -35.4532 | -34.3233
24.1546 | -35.4536 | -34.3289
24.1607 | -35.454 | -34.335
24.1668 | -35.4544 | -34.3408
24.1729 | -35.4548 | -34.3472
24.179 | -35.4553 | -34.3531
24.1851 | -35.4557 | -34.3591
24.1912 | -35.4561 | -34.365
24.1973 | -35.4566 | -34.3707
24.2034 | -35.4571 | -34.376
24.2095 | -35.4575 | -34.3821
24.2156 | -35.458 | -34.3876
24.2217 | -35.4584 | -34.3936
24.2278 | -35.4589 | -34.3991
24.2339 | -35.4593 | -34.4051
24.24 | -35.4596 | -34.4106
24.2461 | -35.46 | -34.4164
24.2522 | -35.4603 | -34.4226
24.2583 | -35.4608 | -34.4283
24.2644 | -35.4612 | -34.434
24.2705 | -35.4617 | -34.4406
24.2766 | -35.462 | -34.4464
24.2827 | -35.4625 | -34.4524
24.2889 | -35.463 | -34.4579
24.295 | -35.4634 | -34.464
24.3011 | -35.4638 | -34.4698
24.3072 | -35.4641 | -34.4755
24.3133 | -35.4646 | -34.4811
24.3194 | -35.465 | -34.4868
24.3255 | -35.4654 | -34.4925
24.3316 | -35.4657 | -34.4984
24.3377 | -35.4662 | -34.5043
24.3438 | -35.4666 | -34.5103
24.3499 | -35.467 | -34.5162
24.356 | -35.4674 | -34.5224
24.3621 | -35.4679 | -34.5279
24.3682 | -35.4683 | -34.534
24.3743 | -35.4687 | -34.5399
24.3804 | -35.469 | -34.5455
24.3865 | -35.4694 | -34.5512
24.3926 | -35.4699 | -34.557
24.3987 | -35.4703 | -34.5633
24.4048 | -35.4707 | -34.569
24.4109 | -35.471 | -34.5748
24.417 | -35.4714 | -34.5806
24.4231 | -35.4717 | -34.5862
24.4293 | -35.4721 | -34.5923
24.4354 | -35.4725 | -34.5984
24.4415 | -35.4729 | -34.604
24.4476 | -35.4733 | -34.6096
24.4537 | -35.4738 | -34.6153
24.4598 | -35.4742 | -34.6209
24.4659 | -35.4746 | -34.6267
24.472 | -35.475 | -34.6326
24.4781 | -35.4755 | -34.6384
24.4842 | -35.4759 | -34.6444
24.4903 | -35.4763 | -34.6502
24.4964 | -35.4767 | -34.656
24.5025 | -35.4772 | -34.6621
24.5086 | -35.4776 | -34.6682
24.5147 | -35.478 | -34.6742
24.5208 | -35.4784 | -34.6806
24.5269 | -35.4788 | -34.6865
24.533 | -35.4792 | -34.6928
24.5391 | -35.4796 | -34.6984
24.5452 | -35.48 | -34.7041
24.5513 | -35.4804 | -34.7103
24.5574 | -35.4807 | -34.7161
24.5635 | -35.4811 | -34.7221
24.5696 | -35.4815 | -34.7282
24.5758 | -35.4819 | -34.7337
24.5819 | -35.4823 | -34.7393
24.588 | -35.4828 | -34.7448
24.5941 | -35.4832 | -34.7505
24.6002 | -35.4836 | -34.7566
24.6063 | -35.484 | -34.7626
24.6124 | -35.4844 | -34.7687
24.6185 | -35.4848 | -34.7744
24.6246 | -35.4852 | -34.7802
24.6307 | -35.4856 | -34.7859
24.6368 | -35.4861 | -34.792
24.6429 | -35.4865 | -34.7978
24.649 | -35.4869 | -34.8039
24.6551 | -35.4873 | -34.8098
24.6612 | -35.4877 | -34.8152
24.6673 | -35.4881 | -34.8212
24.6734 | -35.4885 | -34.8273
24.6795 | -35.4889 | -34.833
24.6856 | -35.4893 | -34.8388
24.6917 | -35.4897 | -34.8448
24.6978 | -35.4901 | -34.8506
24.7039 | -35.4905 | -34.8567
24.71 | -35.4909 | -34.8624
24.7162 | -35.4913 | -34.8685
24.7223 | -35.4917 | -34.8746
24.7284 | -35.4921 | -34.8805
24.7345 | -35.4925 | -34.8863
24.7406 | -35.4929 | -34.892
24.7467 | -35.4933 | -34.8978
24.7528 | -35.4938 | -34.9039
24.7589 | -35.4942 | -34.9095
24.765 | -35.4945 | -34.915
24.7711 | -35.495 | -34.9207
24.7772 | -35.4954 | -34.9266
24.7833 | -35.4959 | -34.9329
24.7894 | -35.4962 | -34.939
24.7955 | -35.4966 | -34.945
24.8016 | -35.497 | -34.9508
24.8077 | -35.4974 | -34.9563
24.8138 | -35.4978 | -34.9623
24.8199 | -35.4982 | -34.9679
24.826 | -35.4986 | -34.9737
24.8321 | -35.4989 | -34.9798
24.8382 | -35.4993 | -34.9855
24.8443 | -35.4997 | -34.9916
24.8504 | -35.5001 | -34.9977
24.8565 | -35.5005 | -35.004
24.8627 | -35.501 | -35.01
24.8688 | -35.5013 | -35.0155
24.8749 | -35.5017 | -35.021
24.881 | -35.502 | -35.0268
24.8871 | -35.5024 | -35.0331
24.8932 | -35.5027 | -35.0389
24.8993 | -35.503 | -35.045
24.9054 | -35.5034 | -35.0506
24.9115 | -35.5037 | -35.0565
24.9176 | -35.5041 | -35.0628
24.9237 | -35.5045 | -35.0687
24.9298 | -35.5049 | -35.0746
24.9359 | -35.5053 | -35.0804
24.942 | -35.5057 | -35.0858
24.9481 | -35.506 | -35.0916
24.9542 | -35.5064 | -35.0974
24.9603 | -35.5068 | -35.1034
24.9664 | -35.5072 | -35.1092
24.9725 | -35.5076 | -35.1155
24.9786 | -35.508 | -35.1213
24.9847 | -35.5084 | -35.1272
24.9908 | -35.5088 | -35.1333
24.9969 | -35.5091 | -35.1392
25.0031 | -35.5095 | -35.1448
25.0092 | -35.5098 | -35.1504
25.0153 | -35.5102 | -35.1562
25.0214 | -35.5106 | -35.1621
25.0275 | -35.511 | -35.1681
25.0336 | -35.5113 | -35.1738
25.0397 | -35.5117 | -35.1796
25.0458 | -35.5121 | -35.1856
25.0519 | -35.5124 | -35.1914
25.058 | -35.5128 | -35.197
25.0641 | -35.5132 | -35.2031
25.0702 | -35.5136 | -35.2088
25.0763 | -35.514 | -35.2149
25.0824 | -35.5143 | -35.2209
25.0885 | -35.5147 | -35.2267
25.0946 | -35.5151 | -35.2326
25.1007 | -35.5155 | -35.2388
25.1068 | -35.5159 | -35.2446
25.1129 | -35.5163 | -35.2504
25.119 | -35.5167 | -35.2563
25.1251 | -35.5171 | -35.2624
25.1312 | -35.5175 | -35.2683
25.1373 | -35.5179 | -35.2744
25.1435 | -35.5183 | -35.2806
25.1496 | -35.5187 | -35.2864
25.1557 | -35.519 | -35.2921
25.1618 | -35.5194 | -35.2979
25.1679 | -35.5198 | -35.3039
25.174 | -35.5201 | -35.31
25.1801 | -35.5205 | -35.3158
25.1862 | -35.5208 | -35.3218
25.1923 | -35.5212 | -35.3276
25.1984 | -35.5215 | -35.3334
25.2045 | -35.5219 | -35.3393
25.2106 | -35.5222 | -35.3449
25.2167 | -35.5226 | -35.3508
25.2228 | -35.5229 | -35.3567
25.2289 | -35.5233 | -35.3627
25.235 | -35.5237 | -35.3685
25.2411 | -35.5241 | -35.3745
25.2472 | -35.5244 | -35.3805
25.2533 | -35.5247 | -35.3864
25.2594 | -35.5251 | -35.392
25.2655 | -35.5255 | -35.398
25.2716 | -35.5258 | -35.4033
25.2777 | -35.5262 | -35.4089
25.2838 | -35.5266 | -35.4148
25.29 | -35.5269 | -35.4205
25.2961 | -35.5273 | -35.4264
25.3022 | -35.5277 | -35.4322
25.3083 | -35.528 | -35.438
25.3144 | -35.5284 | -35.4436
25.3205 | -35.5288 | -35.4494
25.3266 | -35.5291 | -35.4553
25.3327 | -35.5295 | -35.4609
25.3388 | -35.5298 | -35.4667
25.3449 | -35.5302 | -35.4726
25.351 | -35.5306 | -35.4783
25.3571 | -35.5309 | -35.4842
25.3632 | -35.5312 | -35.4902
25.3693 | -35.5315 | -35.4962
25.3754 | -35.5319 | -35.502
25.3815 | -35.5322 | -35.5077
25.3876 | -35.5325 | -35.5136
25.3937 | -35.5329 | -35.5196
25.3998 | -35.5331 | -35.5252
25.4059 | -35.5335 | -35.531
25.412 | -35.5338 | -35.5369
25.4181 | -35.5342 | -35.5429
25.4242 | -35.5345 | -35.549
25.4304 | -35.5349 | -35.5547
25.4365 | -35.5353 | -35.5603
25.4426 | -35.5356 | -35.566
25.4487 | -35.536 | -35.5716
25.4548 | -35.5364 | -35.5776
25.4609 | -35.5367 | -35.5833
25.467 | -35.5371 | -35.5892
25.4731 | -35.5374 | -35.5949
25.4792 | -35.5378 | -35.6006
25.4853 | -35.5382 | -35.6065
25.4914 | -35.5385 | -35.6125
25.4975 | -35.5389 | -35.6182
25.5036 | -35.5392 | -35.624
25.5097 | -35.5396 | -35.6298
25.5158 | -35.54 | -35.6357
25.5219 | -35.5404 | -35.6413
25.528 | -35.5408 | -35.6473
25.5341 | -35.5411 | -35.6531
25.5402 | -35.5415 | -35.659
25.5463 | -35.5418 | -35.6647
25.5524 | -35.5422 | -35.6706
25.5585 | -35.5426 | -35.6763
25.5646 | -35.5429 | -35.6821
25.5707 | -35.5433 | -35.6879
25.5769 | -35.5436 | -35.6934
25.583 | -35.544 | -35.6992
25.5891 | -35.5443 | -35.7049
25.5952 | -35.5447 | -35.7109
25.6013 | -35.545 | -35.7167
25.6074 | -35.5454 | -35.7223
25.6135 | -35.5457 | -35.728
25.6196 | -35.5461 | -35.7339
25.6257 | -35.5464 | -35.7397
25.6318 | -35.5468 | -35.7455
25.6379 | -35.5471 | -35.7515
25.644 | -35.5474 | -35.7572
25.6501 | -35.5478 | -35.7629
25.6562 | -35.5482 | -35.7688
25.6623 | -35.5485 | -35.7745
25.6684 | -35.5489 | -35.7802
25.6745 | -35.5493 | -35.786
25.6806 | -35.5496 | -35.7917
25.6867 | -35.5499 | -35.7976
25.6928 | -35.5503 | -35.8034
25.6989 | -35.5506 | -35.8093
25.705 | -35.5509 | -35.8151
25.7111 | -35.5513 | -35.8209
25.7173 | -35.5516 | -35.8266
25.7234 | -35.552 | -35.8323
25.7295 | -35.5523 | -35.8381
25.7356 | -35.5527 | -35.8439
25.7417 | -35.553 | -35.8495
25.7478 | -35.5533 | -35.8551
25.7539 | -35.5536 | -35.8607
25.76 | -35.554 | -35.8664
25.7661 | -35.5544 | -35.8723
25.7722 | -35.5547 | -35.8782
25.7783 | -35.555 | -35.8841
25.7844 | -35.5553 | -35.8898
25.7905 | -35.5557 | -35.8955
25.7966 | -35.5561 | -35.9011
25.8027 | -35.5564 | -35.907
25.8088 | -35.5567 | -35.9128
25.8149 | -35.557 | -35.9183
25.821 | -35.5573 | -35.9238
25.8271 | -35.5576 | -35.9295
25.8332 | -35.558 | -35.9352
25.8393 | -35.5584 | -35.941
25.8454 | -35.5587 | -35.9469
25.8515 | -35.5591 | -35.9528
25.8576 | -35.5594 | -35.9585
25.8638 | -35.5597 | -35.9641
25.8699 | -35.5601 | -35.9697
25.876 | -35.5604 | -35.9753
25.8821 | -35.5607 | -35.9809
25.8882 | -35.561 | -35.9867
25.8943 | -35.5613 | -35.9925
25.9004 | -35.5616 | -35.9982
25.9065 | -35.562 | -36.0039
25.9126 | -35.5623 | -36.0098
25.9187 | -35.5627 | -36.0155
25.9248 | -35.5629 | -36.0213
25.9309 | -35.5633 | -36.0271
25.937 | -35.5637 | -36.0329
25.9431 | -35.564 | -36.0386
25.9492 | -35.5644 | -36.0441
25.9553 | -35.5647 | -36.0501
25.9614 | -35.565 | -36.0558
25.9675 | -35.5654 | -36.0617
25.9736 | -35.5657 | -36.0675
25.9797 | -35.566 | -36.0734
25.9858 | -35.5664 | -36.0792
25.9919 | -35.5667 | -36.0851
25.998 | -35.567 | -36.0906
26.0042 | -35.5673 | -36.0963
26.0103 | -35.5676 | -36.102
26.0164 | -35.568 | -36.1077
26.0225 | -35.5683 | -36.1134
26.0286 | -35.5687 | -36.1191
26.0347 | -35.5689 | -36.1246
26.0408 | -35.5693 | -36.1304
26.0469 | -35.5696 | -36.1359
26.053 | -35.57 | -36.1415
26.0591 | -35.5703 | -36.1472
26.0652 | -35.5706 | -36.1527
26.0713 | -35.5709 | -36.1584
26.0774 | -35.5712 | -36.1642
26.0835 | -35.5716 | -36.1701
26.0896 | -35.5719 | -36.1758
26.0957 | -35.5722 | -36.1815
26.1018 | -35.5725 | -36.187
26.1079 | -35.5728 | -36.1927
26.114 | -35.5732 | -36.1983
26.1201 | -35.5735 | -36.2039
26.1262 | -35.5739 | -36.2095
26.1323 | -35.5742 | -36.2151
26.1384 | -35.5745 | -36.2205
26.1445 | -35.5749 | -36.2264
26.1507 | -35.5752 | -36.2321
26.1568 | -35.5755 | -36.2376
26.1629 | -35.5759 | -36.2434
26.169 | -35.5762 | -36.249
26.1751 | -35.5765 | -36.2546
26.1812 | -35.5768 | -36.2602
26.1873 | -35.5771 | -36.2661
26.1934 | -35.5775 | -36.2718
26.1995 | -35.5778 | -36.2776
26.2056 | -35.5781 | -36.2831
26.2117 | -35.5784 | -36.2887
26.2178 | -35.5787 | -36.2943
26.2239 | -35.5791 | -36.2997
26.23 | -35.5794 | -36.3051
26.2361 | -35.5797 | -36.3109
26.2422 | -35.58 | -36.3163
26.2483 | -35.5803 | -36.3216
26.2544 | -35.5806 | -36.3271
26.2605 | -35.581 | -36.3326
26.2666 | -35.5813 | -36.3383
26.2727 | -35.5816 | -36.3438
26.2788 | -35.5819 | -36.3491
26.2849 | -35.5822 | -36.3547
26.2911 | -35.5825 | -36.3604
26.2972 | -35.5828 | -36.3661
26.3033 | -35.5832 | -36.3719
26.3094 | -35.5835 | -36.3775
26.3155 | -35.5838 | -36.3828
26.3216 | -35.5841 | -36.3885
26.3277 | -35.5844 | -36.3942
26.3338 | -35.5847 | -36.3997
26.3399 | -35.585 | -36.4052
26.346 | -35.5853 | -36.4104
26.3521 | -35.5856 | -36.4162
26.3582 | -35.5859 | -36.4218
26.3643 | -35.5863 | -36.4274
26.3704 | -35.5866 | -36.4331
26.3765 | -35.5869 | -36.4385
26.3826 | -35.5872 | -36.444
26.3887 | -35.5875 | -36.4496
26.3948 | -35.5878 | -36.4555
26.4009 | -35.5881 | -36.4612
26.407 | -35.5885 | -36.4669
26.4131 | -35.5888 | -36.4727
26.4192 | -35.5891 | -36.478
26.4253 | -35.5895 | -36.4835
26.4314 | -35.5898 | -36.4888
26.4376 | -35.5901 | -36.4945
26.4437 | -35.5904 | -36.5003
26.4498 | -35.5907 | -36.5059
26.4559 | -35.591 | -36.5113
26.462 | -35.5913 | -36.517
26.4681 | -35.5916 | -36.5226
26.4742 | -35.5919 | -36.5281
26.4803 | -35.5922 | -36.5335
26.4864 | -35.5924 | -36.5387
26.4925 | -35.5928 | -36.544
26.4986 | -35.5932 | -36.5499
26.5047 | -35.5935 | -36.5555
26.5108 | -35.5938 | -36.5612
26.5169 | -35.5941 | -36.5668
26.523 | -35.5944 | -36.5719
26.5291 | -35.5947 | -36.5775
26.5352 | -35.5951 | -36.5831
26.5413 | -35.5954 | -36.5888
26.5474 | -35.5957 | -36.5944
26.5535 | -35.596 | -36.6001
26.5596 | -35.5963 | -36.6053
26.5657 | -35.5966 | -36.6109
26.5718 | -35.5969 | -36.6164
26.578 | -35.5972 | -36.622
26.5841 | -35.5975 | -36.6275
26.5902 | -35.5978 | -36.6331
26.5963 | -35.5981 | -36.6382
26.6024 | -35.5984 | -36.6438
26.6085 | -35.5987 | -36.6494
26.6146 | -35.599 | -36.655
26.6207 | -35.5993 | -36.6606
26.6268 | -35.5996 | -36.6661
26.6329 | -35.5999 | -36.6714
26.639 | -35.6002 | -36.6768
26.6451 | -35.6005 | -36.6823
26.6512 | -35.6008 | -36.6877
26.6573 | -35.6012 | -36.6932
26.6634 | -35.6014 | -36.6984
26.6695 | -35.6017 | -36.7041
26.6756 | -35.6021 | -36.7099
26.6817 | -35.6024 | -36.7152
26.6878 | -35.6026 | -36.7209
26.6939 | -35.603 | -36.7264
26.7 | -35.6032 | -36.7317
26.7061 | -35.6035 | -36.7371
26.7122 | -35.6039 | -36.7425
26.7183 | -35.6042 | -36.7481
26.7245 | -35.6045 | -36.7536
26.7306 | -35.6048 | -36.7592
26.7367 | -35.6051 | -36.7643
26.7428 | -35.6054 | -36.7696
26.7489 | -35.6057 | -36.7752
26.755 | -35.6061 | -36.7805
26.7611 | -35.6064 | -36.7861
26.7672 | -35.6067 | -36.7915
26.7733 | -35.6069 | -36.7966
26.7794 | -35.6072 | -36.8022
26.7855 | -35.6075 | -36.8076
26.7916 | -35.6078 | -36.813
26.7977 | -35.6082 | -36.8186
26.8038 | -35.6085 | -36.8238
26.8099 | -35.6088 | -36.8292
26.816 | -35.6091 | -36.8347
26.8221 | -35.6094 | -36.8402
26.8282 | -35.6097 | -36.8456
26.8343 | -35.61 | -36.8511
26.8404 | -35.6102 | -36.8563
26.8465 | -35.6106 | -36.8618
26.8526 | -35.6109 | -36.8674
26.8587 | -35.6112 | -36.8728
26.8649 | -35.6115 | -36.8781
26.871 | -35.6118 | -36.8836
26.8771 | -35.6121 | -36.8888
26.8832 | -35.6124 | -36.8945
26.8893 | -35.6127 | -36.9
26.8954 | -35.613 | -36.9052
26.9015 | -35.6133 | -36.9107
26.9076 | -35.6136 | -36.9163
26.9137 | -35.6139 | -36.9217
26.9198 | -35.6142 | -36.9272
26.9259 | -35.6145 | -36.9325
26.932 | -35.6148 | -36.9379
26.9381 | -35.6151 | -36.9433
26.9442 | -35.6153 | -36.9484
26.9503 | -35.6156 | -36.9535
26.9564 | -35.6159 | -36.9587
26.9625 | -35.6163 | -36.9643
26.9686 | -35.6165 | -36.9698
26.9747 | -35.6168 | -36.9751
26.9808 | -35.617 | -36.9802
26.9869 | -35.6173 | -36.9855
26.993 | -35.6176 | -36.9909
26.9991 | -35.6179 | -36.9964
27.0052 | -35.6182 | -37.0019
27.0114 | -35.6185 | -37.0074
27.0175 | -35.6187 | -37.0124
27.0236 | -35.619 | -37.0179
27.0297 | -35.6193 | -37.0233
27.0358 | -35.6196 | -37.0287
27.0419 | -35.6199 | -37.0339
27.048 | -35.6202 | -37.0394
27.0541 | -35.6205 | -37.0446
27.0602 | -35.6208 | -37.05
27.0663 | -35.6211 | -37.0556
27.0724 | -35.6214 | -37.0608
27.0785 | -35.6217 | -37.0661
27.0846 | -35.622 | -37.0711
27.0907 | -35.6223 | -37.0766
27.0968 | -35.6226 | -37.0819
27.1029 | -35.6229 | -37.0873
27.109 | -35.6232 | -37.0928
27.1151 | -35.6235 | -37.0982
27.1212 | -35.6237 | -37.1034
27.1273 | -35.6241 | -37.1087
27.1334 | -35.6243 | -37.114
27.1395 | -35.6246 | -37.1196
27.1456 | -35.6249 | -37.1249
27.1518 | -35.6252 | -37.1302
27.1579 | -35.6255 | -37.1352
27.164 | -35.6258 | -37.1404
27.1701 | -35.6261 | -37.1456
27.1762 | -35.6264 | -37.1508
27.1823 | -35.6266 | -37.1562
27.1884 | -35.6269 | -37.1615
27.1945 | -35.6272 | -37.1662
27.2006 | -35.6275 | -37.1717
27.2067 | -35.6278 | -37.177
27.2128 | -35.628 | -37.1824
27.2189 | -35.6283 | -37.1876
27.225 | -35.6286 | -37.193
27.2311 | -35.6289 | -37.1979
27.2372 | -35.6292 | -37.2032
27.2433 | -35.6295 | -37.2085
27.2494 | -35.6298 | -37.2136
27.2555 | -35.6301 | -37.219
27.2616 | -35.6304 | -37.2241
27.2677 | -35.6307 | -37.2292
27.2738 | -35.631 | -37.2344
27.2799 | -35.6312 | -37.2397
27.286 | -35.6316 | -37.2449
27.2921 | -35.6319 | -37.2503
27.2983 | -35.6322 | -37.2552
27.3044 | -35.6325 | -37.2607
27.3105 | -35.6328 | -37.266
27.3166 | -35.6331 | -37.2711
27.3227 | -35.6333 | -37.2763
27.3288 | -35.6337 | -37.2818
27.3349 | -35.6339 | -37.2868
27.341 | -35.6342 | -37.2922
27.3471 | -35.6345 | -37.2973
27.3532 | -35.6348 | -37.3023
27.3593 | -35.6351 | -37.3077
27.3654 | -35.6354 | -37.313
27.3715 | -35.6357 | -37.3179
27.3776 | -35.636 | -37.3233
27.3837 | -35.6363 | -37.3283
27.3898 | -35.6366 | -37.3336
27.3959 | -35.6369 | -37.339
27.402 | -35.6371 | -37.3438
27.4081 | -35.6375 | -37.3491
27.4142 | -35.6378 | -37.3543
27.4203 | -35.6381 | -37.3596
27.4264 | -35.6384 | -37.3651
27.4325 | -35.6387 | -37.37
27.4387 | -35.6389 | -37.3749
27.4448 | -35.6392 | -37.38
27.4509 | -35.6395 | -37.3853
27.457 | -35.6397 | -37.3907
27.4631 | -35.64 | -37.396
27.4692 | -35.6404 | -37.4009
27.4753 | -35.6406 | -37.4057
27.4814 | -35.6409 | -37.4107
27.4875 | -35.6412 | -37.4158
27.4936 | -35.6415 | -37.4209
27.4997 | -35.6417 | -37.4262
27.5058 | -35.642 | -37.4315
27.5119 | -35.6423 | -37.4362
27.518 | -35.6426 | -37.4413
27.5241 | -35.6428 | -37.4466
27.5302 | -35.6431 | -37.4521
27.5363 | -35.6434 | -37.4572
27.5424 | -35.6437 | -37.4621
27.5485 | -35.644 | -37.4674
27.5546 | -35.6443 | -37.4722
27.5607 | -35.6446 | -37.4772
27.5668 | -35.6449 | -37.4823
27.5729 | -35.6451 | -37.4876
27.5791 | -35.6454 | -37.4926
27.5852 | -35.6457 | -37.4977
27.5913 | -35.6459 | -37.5026
27.5974 | -35.6463 | -37.5079
27.6035 | -35.6465 | -37.513
27.6096 | -35.6468 | -37.5182
27.6157 | -35.6471 | -37.5229
27.6218 | -35.6474 | -37.5282
27.6279 | -35.6477 | -37.5334
27.634 | -35.6479 | -37.5387
27.6401 | -35.6482 | -37.5439
27.6462 | -35.6485 | -37.5489
27.6523 | -35.6488 | -37.5537
27.6584 | -35.6491 | -37.559
27.6645 | -35.6493 | -37.564
27.6706 | -35.6496 | -37.5692
27.6767 | -35.65 | -37.5743
27.6828 | -35.6502 | -37.5791
27.6889 | -35.6505 | -37.5843
27.695 | -35.6508 | -37.5894
27.7011 | -35.651 | -37.5943
27.7072 | -35.6513 | -37.5994
27.7133 | -35.6516 | -37.6049
27.7194 | -35.6519 | -37.6098
27.7256 | -35.6521 | -37.615
27.7317 | -35.6524 | -37.6199
27.7378 | -35.6527 | -37.6251
27.7439 | -35.6529 | -37.6304
27.75 | -35.6532 | -37.6357
27.7561 | -35.6535 | -37.6406
27.7622 | -35.6538 | -37.6457
27.7683 | -35.6541 | -37.6509
27.7744 | -35.6544 | -37.656
27.7805 | -35.6546 | -37.6611
27.7866 | -35.6549 | -37.6661
27.7927 | -35.6552 | -37.6708
27.7988 | -35.6555 | -37.6759
27.8049 | -35.6558 | -37.6809
27.811 | -35.6561 | -37.6859
27.8171 | -35.6564 | -37.6908
27.8232 | -35.6567 | -37.6954
27.8293 | -35.6569 | -37.7008
27.8354 | -35.6572 | -37.706
27.8415 | -35.6575 | -37.7111
27.8476 | -35.6578 | -37.7161
27.8537 | -35.658 | -37.7213
27.8598 | -35.6583 | -37.7263
27.866 | -35.6586 | -37.7313
27.8721 | -35.6588 | -37.7366
27.8782 | -35.6591 | -37.7418
27.8843 | -35.6593 | -37.7467
27.8904 | -35.6596 | -37.7518
27.8965 | -35.6599 | -37.7566
27.9026 | -35.6601 | -37.7619
27.9087 | -35.6605 | -37.7672
27.9148 | -35.6608 | -37.7722
27.9209 | -35.661 | -37.7773
27.927 | -35.6613 | -37.7823
27.9331 | -35.6616 | -37.7871
27.9392 | -35.6619 | -37.7923
27.9453 | -35.6622 | -37.7974
27.9514 | -35.6624 | -37.8024
27.9575 | -35.6627 | -37.8073
27.9636 | -35.663 | -37.8125
27.9697 | -35.6632 | -37.8172
27.9758 | -35.6635 | -37.8226
27.9819 | -35.6637 | -37.8277
27.988 | -35.664 | -37.8326
27.9941 | -35.6642 | -37.8375
28.0002 | -35.6645 | -37.8422
28.0063 | -35.6647 | -37.847
28.0125 | -35.665 | -37.8522
28.0186 | -35.6652 | -37.8571
28.0247 | -35.6655 | -37.8622
28.0308 | -35.6658 | -37.867
28.0369 | -35.6661 | -37.8717
28.043 | -35.6664 | -37.877
28.0491 | -35.6667 | -37.8819
28.0552 | -35.6669 | -37.8868
28.0613 | -35.6672 | -37.8918
28.0674 | -35.6675 | -37.8969
28.0735 | -35.6678 | -37.9016

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4944 | -42.6733
21.4908 | -35.5244 | -34.9435
16.4948 | -34.9059 | -27.6156
12.6602 | -34.5396 | -21.3786
9.7171 | -34.3256 | -16.4924
7.4582 | -34.1742 | -12.5276
5.7244 | -34.0691 | -9.3228
4.3936 | -33.998 | -6.7221
3.3722 | -33.9785 | -4.6751
2.5883 | -33.9656 | -3.1275
1.9866 | -33.9618 | -1.8901
1.5248 | -33.9643 | -0.8998
1.1703 | -33.9601 | -0.25251
1 | -33.9641 | 0.066933
0.89822 | -33.9509 | 0.33628
0.68941 | -33.9447 | 0.91581
0.52919 | -33.9513 | 1.3633
0.40612 | -33.9375 | 1.9722
0.31171 | -33.9305 | 2.4315
0.23928 | -33.9333 | 3.1588
0.18358 | -33.9447 | 3.5895
0.14091 | -33.9655 | 4.4711
0.10819 | -33.9712 | 5.6471
0.083017 | -34.0179 | 7.1249
0.063705 | -34.057 | 9.0079
0.048923 | -34.1215 | 11.5087
0.037479 | -34.2247 | 14.7072
0.028801 | -34.3973 | 18.732
0.022078 | -34.6935 | 24.0965
0.017023 | -35.169 | 30.5117
0.013018 | -36.0087 | 38.2077
0.010014 | -37.118 | 46.0999

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003704 | -9.6622
0.007408 | -27.3875
0.011112 | -36.1938
0.014816 | -41.1145
0.018519 | -43.6405
0.022223 | -45.568
0.025927 | -46.916
0.029631 | -48.3311
0.033335 | -49.9526
0.037039 | -50.2594
0.040743 | -50.9588
0.044447 | -52.5838
0.04815 | -53.4818
0.051854 | -55.0077
0.055558 | -56.3115
0.059262 | -56.8564
0.062966 | -56.6976
0.06667 | -57.8079
0.070374 | -58.7742
0.074078 | -59.2511
0.077781 | -60.4127
0.081485 | -60.5973
0.085189 | -60.4514
0.088893 | -60.7803
0.092597 | -61.6663
0.096301 | -62.5726
0.1 | -63.0009
0.10371 | -62.927
0.10741 | -63.2242
0.11112 | -63.7559
0.11482 | -63.66
0.11852 | -65.4719
0.12223 | -67.1191
0.12593 | -67.6235
0.12964 | -67.1275
0.13334 | -67.1111
0.13704 | -67.1118
0.14075 | -67.3367
0.14445 | -68.4394
0.14816 | -69.6744
0.15186 | -67.6663
0.15556 | -68.2647
0.15927 | -68.5253
0.16297 | -68.8463
0.16667 | -68.2227
0.17038 | -68.5256
0.17408 | -69.2688
0.17779 | -69.813
0.18149 | -71.2452
0.18519 | -71.9165
0.1889 | -71.337
0.1926 | -71.1373
0.19631 | -71.5332
0.20001 | -72.551
0.20371 | -72.1686
0.20742 | -70.9811
0.21112 | -70.9114
0.21482 | -71.3763
0.21853 | -71.5858
0.22223 | -72.467
0.22594 | -73.8378
0.22964 | -74.1605
0.23334 | -73.4801
0.23705 | -73.0052
0.24075 | -72.5344
0.24446 | -72.7922
0.24816 | -73.0336
0.25186 | -72.7439
0.25557 | -72.961
0.25927 | -74.3787
0.26298 | -74.5316
0.26668 | -74.6484
0.27038 | -74.4505
0.27409 | -74.1094
0.27779 | -74.286
0.28149 | -74.2444
0.2852 | -72.6905
0.2889 | -73.6364
0.29261 | -74.7418
0.29631 | -75.478
0.30001 | -76.0993
0.30372 | -76.8744
0.30742 | -76.6233
0.31113 | -76.663
0.31483 | -76.0119
0.31853 | -75.7966
0.32224 | -77.0418
0.32594 | -76.9417
0.32965 | -76.1647
0.33335 | -76.627
0.33705 | -76.2917
0.34076 | -77.0033
0.34446 | -77.5821
0.34816 | -78.2186
0.35187 | -78.7989
0.35557 | -78.5577
0.35928 | -78.5912
0.36298 | -78.9301
0.36668 | -78.2043
0.37039 | -78.0071
0.37409 | -78.2854
0.37779 | -78.6353
0.3815 | -78.6673
0.3852 | -77.8295
0.38891 | -78.7135
0.39261 | -78.9435
0.39631 | -78.3993
0.40002 | -78.4592
0.40372 | -79.4789
0.40743 | -79.3477
0.41113 | -79.4275
0.41483 | -78.9993
0.41854 | -78.8558
0.42224 | -78.4083
0.42595 | -78.5844
0.42965 | -79.7711
0.43335 | -79.5186
0.43706 | -79.6627
0.44076 | -80.1349
0.44446 | -81.1205
0.44817 | -81.4098
0.45187 | -80.5715
0.45558 | -79.4487
0.45928 | -79.5343
0.46298 | -80.1771
0.46669 | -80.0613
0.47039 | -79.4737
0.4741 | -79.8611
0.4778 | -79.4184
0.4815 | -78.6443
0.48521 | -79.2477
0.48891 | -79.9445
0.49262 | -80.22
0.49632 | -79.8793
0.50002 | -79.9802
0.50373 | -80.7742
0.50743 | -80.9277
0.51114 | -80.4141
0.51484 | -79.7822
0.51854 | -80.1054
0.52225 | -80.2739
0.52595 | -81.4459
0.52965 | -82.6353
0.53336 | -81.7932
0.53706 | -81.3722
0.54077 | -81.2652
0.54447 | -81.4529
0.54817 | -81.7807
0.55188 | -81.7956
0.55558 | -82.0676
0.55929 | -82.504
0.56299 | -82.5656
0.56669 | -82.4472
0.5704 | -82.3874
0.5741 | -82.102
0.57781 | -81.5762
0.58151 | -81.4637
0.58521 | -81.9667
0.58892 | -82.6548
0.59262 | -82.9644
0.59632 | -83.0032
0.60003 | -83.0422
0.60373 | -82.8871
0.60744 | -82.3393
0.61114 | -81.961
0.61484 | -82.3811
0.61855 | -82.9144
0.62225 | -83.1018
0.62596 | -83.3092
0.62966 | -83.4522
0.63336 | -83.6336
0.63707 | -83.781
0.64077 | -83.6956
0.64447 | -83.3805
0.64818 | -83.0984
0.65188 | -83.023
0.65559 | -83.399
0.65929 | -83.4407
0.66299 | -83.3059
0.6667 | -83.5055
0.6704 | -84.0339
0.6741 | -84.5077
0.67781 | -84.3426
0.68151 | -83.9297
0.68522 | -84.0793
0.68892 | -84.2761
0.69263 | -84.243
0.69633 | -84.4628
0.70003 | -84.5384
0.70374 | -84.5073
0.70744 | -84.51
0.71114 | -84.4036
0.71485 | -84.3024
0.71855 | -84.6378
0.72226 | -84.8915
0.72596 | -84.6122
0.72966 | -84.0978
0.73337 | -83.6842
0.73707 | -83.7107
0.74077 | -83.846
0.74448 | -83.8817
0.74818 | -84.0025
0.75189 | -84.1795
0.75559 | -84.225
0.75929 | -84.276
0.763 | -84.3993
0.7667 | -84.7349
0.77041 | -84.9257
0.77411 | -84.5761
0.77781 | -84.3709
0.78152 | -84.4576
0.78522 | -84.7552
0.78892 | -85.4592
0.79263 | -86.3284
0.79633 | -86.6195
0.80004 | -86.3156
0.80374 | -85.7481
0.80744 | -85.0613
0.81115 | -84.7757
0.81485 | -85.1487
0.81856 | -85.7497
0.82226 | -85.886
0.82596 | -85.3737
0.82967 | -84.9599
0.83337 | -85.0476
0.83708 | -85.1464
0.84078 | -84.8268
0.84448 | -84.9165
0.84819 | -85.5869
0.85189 | -86.5678
0.85559 | -87.6656
0.8593 | -87.7269
0.863 | -86.5604
0.86671 | -85.7733
0.87041 | -85.7989
0.87411 | -86.056
0.87782 | -86.0811
0.88152 | -85.9903
0.88523 | -86.2912
0.88893 | -86.537
0.89263 | -86.3016
0.89634 | -86.1791
0.90004 | -86.4917
0.90375 | -86.9295
0.90745 | -87.1097
0.91115 | -86.7122
0.91486 | -86.2501
0.91856 | -85.9594
0.92227 | -85.7774
0.92597 | -85.8694
0.92967 | -85.7704
0.93338 | -85.8648
0.93708 | -86.507
0.94078 | -87.1136
0.94449 | -87.3049
0.94819 | -87.3084
0.9519 | -87.1745
0.9556 | -87.2367
0.9593 | -87.362
0.96301 | -87.4234
0.96671 | -87.3006
0.97042 | -87.0207
0.97412 | -86.9158
0.97782 | -86.9349
0.98153 | -86.9842
0.98523 | -86.97
0.98894 | -86.7302
0.99264 | -86.5728
0.99634 | -86.9035
1 | -87.4798
1.0037 | -87.5454
1.0075 | -87.1492
1.0112 | -86.636
1.0149 | -86.358
1.0186 | -86.5272
1.0223 | -87.1178
1.026 | -87.7473
1.0297 | -88.0219
1.0334 | -87.6479
1.0371 | -87.3699
1.0408 | -87.5884
1.0445 | -88.1712
1.0482 | -88.351
1.0519 | -88.1264
1.0556 | -88.0246
1.0593 | -87.9175
1.063 | -87.9186
1.0667 | -87.8666
1.0704 | -87.8328
1.0741 | -87.9145
1.0778 | -88.0428
1.0815 | -88.1347
1.0852 | -88.2818
1.0889 | -88.4855
1.0926 | -88.5325
1.0963 | -88.3465
1.1001 | -88.1381
1.1038 | -88.0445
1.1075 | -88.0307
1.1112 | -87.9838
1.1149 | -87.8867
1.1186 | -87.9294
1.1223 | -88.2323
1.126 | -88.66
1.1297 | -88.9592
1.1334 | -88.9813
1.1371 | -88.7179
1.1408 | -88.311
1.1445 | -88.0007
1.1482 | -87.9603
1.1519 | -88.2473
1.1556 | -88.666
1.1593 | -88.863
1.163 | -88.8812
1.1667 | -89.0038
1.1704 | -89.1699
1.1741 | -89.1446
1.1778 | -88.9214
1.1815 | -88.5796
1.1852 | -88.247
1.1889 | -88.104
1.1926 | -88.2075
1.1964 | -88.4463
1.2001 | -88.5924
1.2038 | -88.4536
1.2075 | -88.2062
1.2112 | -88.1981
1.2149 | -88.5279
1.2186 | -88.919
1.2223 | -89.0498
1.226 | -89.0705
1.2297 | -89.1803
1.2334 | -89.2415
1.2371 | -89.2346
1.2408 | -89.2161
1.2445 | -89.1271
1.2482 | -88.9945
1.2519 | -88.8796
1.2556 | -88.8745
1.2593 | -89.0432
1.263 | -89.3097
1.2667 | -89.592
1.2704 | -89.8876
1.2741 | -90.1815
1.2778 | -90.295
1.2815 | -89.9814
1.2852 | -89.4089
1.2889 | -89.0077
1.2927 | -88.9143
1.2964 | -89.0504
1.3001 | -89.3914
1.3038 | -89.7755
1.3075 | -89.8209
1.3112 | -89.4467
1.3149 | -88.98
1.3186 | -88.6436
1.3223 | -88.4919
1.326 | -88.6055
1.3297 | -89.0016
1.3334 | -89.4431
1.3371 | -89.6546
1.3408 | -89.6619
1.3445 | -89.5651
1.3482 | -89.403
1.3519 | -89.3411
1.3556 | -89.4948
1.3593 | -89.6352
1.363 | -89.4721
1.3667 | -89.2464
1.3704 | -89.2305
1.3741 | -89.326
1.3778 | -89.5084
1.3815 | -89.8346
1.3852 | -90.0923
1.389 | -90.0668
1.3927 | -89.9786
1.3964 | -90.2059
1.4001 | -90.8033
1.4038 | -91.3878
1.4075 | -91.5062
1.4112 | -91.1447
1.4149 | -90.5767
1.4186 | -90.1678
1.4223 | -90.1492
1.426 | -90.4628
1.4297 | -90.9333
1.4334 | -91.3919
1.4371 | -91.6791
1.4408 | -91.7097
1.4445 | -91.5612
1.4482 | -91.3501
1.4519 | -91.0754
1.4556 | -90.8544
1.4593 | -90.8937
1.463 | -91.056
1.4667 | -90.9691
1.4704 | -90.7196
1.4741 | -90.7064
1.4778 | -90.9163
1.4815 | -91.1212
1.4853 | -91.3484
1.489 | -91.5865
1.4927 | -91.5755
1.4964 | -91.22
1.5001 | -90.8617
1.5038 | -90.8387
1.5075 | -91.1049
1.5112 | -91.2138
1.5149 | -90.8447
1.5186 | -90.3769
1.5223 | -90.4233
1.526 | -91.1541
1.5297 | -92.0965
1.5334 | -92.4915
1.5371 | -92.3436
1.5408 | -92.1643
1.5445 | -92.213
1.5482 | -92.3595
1.5519 | -92.3075
1.5556 | -91.9748
1.5593 | -91.5327
1.563 | -91.1655
1.5667 | -90.9452
1.5704 | -90.9355
1.5741 | -91.0961
1.5779 | -91.2874
1.5816 | -91.4727
1.5853 | -91.6228
1.589 | -91.7615
1.5927 | -92.0483
1.5964 | -92.4322
1.6001 | -92.3991
1.6038 | -92.2769
1.6075 | -92.0291
1.6112 | -91.7621
1.6149 | -91.5435
1.6186 | -91.423
1.6223 | -91.4453
1.626 | -91.5977
1.6297 | -91.7819
1.6334 | -91.9116
1.6371 | -91.9434
1.6408 | -91.9068
1.6445 | -91.805
1.6482 | -91.632
1.6519 | -91.4444
1.6556 | -91.3534
1.6593 | -91.3962
1.663 | -91.5108
1.6667 | -91.6439
1.6704 | -91.8208
1.6742 | -92.0071
1.6779 | -92.0714
1.6816 | -91.949
1.6853 | -91.7487
1.689 | -91.6062
1.6927 | -91.5777
1.6964 | -91.6704
1.7001 | -91.8624
1.7038 | -92.0968
1.7075 | -92.3189
1.7112 | -92.5063
1.7149 | -92.6607
1.7186 | -92.7506
1.7223 | -92.7235
1.726 | -92.5796
1.7297 | -92.3859
1.7334 | -92.2077
1.7371 | -92.0565
1.7408 | -91.9595
1.7445 | -91.9561
1.7482 | -92.0343
1.7519 | -92.1483
1.7556 | -92.266
1.7593 | -92.3978
1.763 | -92.5237
1.7667 | -92.6192
1.7705 | -92.6712
1.7742 | -92.693
1.7779 | -92.7071
1.7816 | -92.7251
1.7853 | -92.7386
1.789 | -92.765
1.7927 | -92.8107
1.7964 | -92.848
1.8001 | -92.8486
1.8038 | -92.8495
1.8075 | -92.8449
1.8112 | -92.78
1.8149 | -92.6048
1.8186 | -92.3392
1.8223 | -92.0429
1.826 | -91.7806
1.8297 | -91.6337
1.8334 | -91.6722
1.8371 | -91.9552
1.8408 | -92.465
1.8445 | -93.0407
1.8482 | -93.4513
1.8519 | -93.5948
1.8556 | -93.5464
1.8593 | -93.4399
1.863 | -93.3535
1.8668 | -93.3297
1.8705 | -93.3181
1.8742 | -93.2694
1.8779 | -93.2175
1.8816 | -93.2164
1.8853 | -93.2404
1.889 | -93.1605
1.8927 | -92.9332
1.8964 | -92.6255
1.9001 | -92.3458
1.9038 | -92.1996
1.9075 | -92.2412
1.9112 | -92.4286
1.9149 | -92.6603
1.9186 | -92.8669
1.9223 | -93.0692
1.926 | -93.2643
1.9297 | -93.4481
1.9334 | -93.635
1.9371 | -93.8113
1.9408 | -93.858
1.9445 | -93.7022
1.9482 | -93.4433
1.9519 | -93.2188
1.9556 | -93.0862
1.9593 | -93.0509
1.9631 | -93.1189
1.9668 | -93.2819
1.9705 | -93.5119
1.9742 | -93.733
1.9779 | -93.8418
1.9816 | -93.8246
1.9853 | -93.7374
1.989 | -93.6305
1.9927 | -93.5567
1.9964 | -93.566
2.0001 | -93.6578
2.0038 | -93.7576
2.0075 | -93.7928
2.0112 | -93.7546
2.0149 | -93.6823
2.0186 | -93.641
2.0223 | -93.6716
2.026 | -93.7713
2.0297 | -93.8795
2.0334 | -93.9272
2.0371 | -93.8497
2.0408 | -93.6112
2.0445 | -93.2801
2.0482 | -93.0413
2.0519 | -93.0179
2.0557 | -93.1526
2.0594 | -93.2834
2.0631 | -93.3117
2.0668 | -93.2623
2.0705 | -93.2402
2.0742 | -93.3135
2.0779 | -93.4432
2.0816 | -93.4991
2.0853 | -93.3835
2.089 | -93.0808
2.0927 | -92.7414
2.0964 | -92.5562
2.1001 | -92.6314
2.1038 | -92.918
2.1075 | -93.2545
2.1112 | -93.5159
2.1149 | -93.6936
2.1186 | -93.8248
2.1223 | -93.9083
2.126 | -93.9236
2.1297 | -93.9169
2.1334 | -93.8996
2.1371 | -93.9641
2.1408 | -94.0743
2.1445 | -94.2633
2.1482 | -94.4681
2.152 | -94.578
2.1557 | -94.4925
2.1594 | -94.2838
2.1631 | -94.0706
2.1668 | -93.9424
2.1705 | -93.9322
2.1742 | -94.0269
2.1779 | -94.1976
2.1816 | -94.4005
2.1853 | -94.5987
2.189 | -94.7389
2.1927 | -94.7785
2.1964 | -94.7017
2.2001 | -94.4988
2.2038 | -94.232
2.2075 | -93.9899
2.2112 | -93.873
2.2149 | -93.8982
2.2186 | -94.0456
2.2223 | -94.2539
2.226 | -94.4552
2.2297 | -94.6259
2.2334 | -94.7512
2.2371 | -94.8174
2.2408 | -94.8266
2.2445 | -94.7859
2.2483 | -94.6936
2.252 | -94.5462
2.2557 | -94.3889
2.2594 | -94.2484
2.2631 | -94.1775
2.2668 | -94.2027
2.2705 | -94.2984
2.2742 | -94.4066
2.2779 | -94.4739
2.2816 | -94.4667
2.2853 | -94.4292
2.289 | -94.4118
2.2927 | -94.4055
2.2964 | -94.3587
2.3001 | -94.2646
2.3038 | -94.1797
2.3075 | -94.1547
2.3112 | -94.1876
2.3149 | -94.2266
2.3186 | -94.2394
2.3223 | -94.2526
2.326 | -94.289
2.3297 | -94.3587
2.3334 | -94.4427
2.3371 | -94.5132
2.3408 | -94.5704
2.3446 | -94.6615
2.3483 | -94.8365
2.352 | -95.0891
2.3557 | -95.3434
2.3594 | -95.462
2.3631 | -95.3571
2.3668 | -95.0907
2.3705 | -94.8016
2.3742 | -94.6182
2.3779 | -94.605
2.3816 | -94.7443
2.3853 | -94.9634
2.389 | -95.1724
2.3927 | -95.2782
2.3964 | -95.244
2.4001 | -95.1081
2.4038 | -94.9419
2.4075 | -94.8156
2.4112 | -94.7663
2.4149 | -94.7841
2.4186 | -94.8311
2.4223 | -94.8701
2.426 | -94.8864
2.4297 | -94.8876
2.4334 | -94.9154
2.4371 | -94.9768
2.4409 | -95.0457
2.4446 | -95.0899
2.4483 | -95.0783
2.452 | -95.0174
2.4557 | -94.9184
2.4594 | -94.8271
2.4631 | -94.804
2.4668 | -94.8891
2.4705 | -95.0658
2.4742 | -95.2849
2.4779 | -95.484
2.4816 | -95.591
2.4853 | -95.5462
2.489 | -95.3839
2.4927 | -95.1741
2.4964 | -95.0068
2.5001 | -94.9157
2.5038 | -94.8881
2.5075 | -94.9065
2.5112 | -94.9621
2.5149 | -95.0419
2.5186 | -95.1471
2.5223 | -95.2962
2.526 | -95.5058
2.5297 | -95.7536
2.5335 | -95.9694
2.5372 | -96.0665
2.5409 | -96.0061
2.5446 | -95.8479
2.5483 | -95.6761
2.552 | -95.5641
2.5557 | -95.5467
2.5594 | -95.6406
2.5631 | -95.8244
2.5668 | -96.0585
2.5705 | -96.2768
2.5742 | -96.4246
2.5779 | -96.4392
2.5816 | -96.3245
2.5853 | -96.1072
2.589 | -95.859
2.5927 | -95.6424
2.5964 | -95.4987
2.6001 | -95.4478
2.6038 | -95.5081
2.6075 | -95.6762
2.6112 | -95.9066
2.6149 | -96.0823
2.6186 | -96.1219
2.6223 | -96.0308
2.626 | -95.9115
2.6298 | -95.8405
2.6335 | -95.8457
2.6372 | -95.9043
2.6409 | -95.9607
2.6446 | -95.9356
2.6483 | -95.7804
2.652 | -95.521
2.6557 | -95.232
2.6594 | -95.0171
2.6631 | -94.9443
2.6668 | -95.0295
2.6705 | -95.2524
2.6742 | -95.4165
2.6779 | -95.5686
2.6816 | -95.677
2.6853 | -95.7472
2.689 | -95.7915
2.6927 | -95.8159
2.6964 | -95.8207
2.7001 | -95.8241
2.7038 | -95.8218
2.7075 | -95.8159
2.7112 | -95.8133
2.7149 | -95.8382
2.7186 | -95.8811
2.7223 | -95.9263
2.7261 | -95.9763
2.7298 | -96.039
2.7335 | -96.1294
2.7372 | -96.2716
2.7409 | -96.4442
2.7446 | -96.6057
2.7483 | -96.7157
2.752 | -96.7535
2.7557 | -96.7007
2.7594 | -96.5722
2.7631 | -96.3857
2.7668 | -96.2059
2.7705 | -96.0659
2.7742 | -95.9836
2.7779 | -95.9809
2.7816 | -96.0313
2.7853 | -96.1016
2.789 | -96.158
2.7927 | -96.17
2.7964 | -96.1167
2.8001 | -96.0316
2.8038 | -95.9543
2.8075 | -95.9016
2.8112 | -95.8932
2.8149 | -95.9404
2.8186 | -96.0481
2.8224 | -96.1986
2.8261 | -96.3527
2.8298 | -96.469
2.8335 | -96.503
2.8372 | -96.4594
2.8409 | -96.3597
2.8446 | -96.2375
2.8483 | -96.1293
2.852 | -96.0742
2.8557 | -96.0787
2.8594 | -96.1371
2.8631 | -96.2046
2.8668 | -96.2363
2.8705 | -96.2163
2.8742 | -96.176
2.8779 | -96.1638
2.8816 | -96.2129
2.8853 | -96.303
2.889 | -96.3982
2.8927 | -96.4658
2.8964 | -96.4901
2.9001 | -96.5178
2.9038 | -96.5748
2.9075 | -96.6588
2.9112 | -96.754
2.915 | -96.8264
2.9187 | -96.839
2.9224 | -96.7923
2.9261 | -96.7018
2.9298 | -96.5837
2.9335 | -96.4673
2.9372 | -96.3502
2.9409 | -96.2441
2.9446 | -96.1732
2.9483 | -96.1419
2.952 | -96.1851
2.9557 | -96.293
2.9594 | -96.4401
2.9631 | -96.5856
2.9668 | -96.7024
2.9705 | -96.7713
2.9742 | -96.8151
2.9779 | -96.8585
2.9816 | -96.9219
2.9853 | -97.0006
2.989 | -97.0449
2.9927 | -97.0364
2.9964 | -96.971
3.0001 | -96.8366
3.0038 | -96.6651
3.0075 | -96.5144
3.0113 | -96.433
3.015 | -96.4384
3.0187 | -96.5133
3.0224 | -96.6206
3.0261 | -96.71
3.0298 | -96.7413
3.0335 | -96.7099
3.0372 | -96.6428
3.0409 | -96.5671
3.0446 | -96.5179
3.0483 | -96.5175
3.052 | -96.5708
3.0557 | -96.6516
3.0594 | -96.7319
3.0631 | -96.7804
3.0668 | -96.7618
3.0705 | -96.6691
3.0742 | -96.5295
3.0779 | -96.3995
3.0816 | -96.3285
3.0853 | -96.3347
3.089 | -96.4115
3.0927 | -96.54
3.0964 | -96.6849
3.1001 | -96.8175
3.1038 | -96.9237
3.1076 | -96.9922
3.1113 | -97.0325
3.115 | -97.0525
3.1187 | -97.0596
3.1224 | -97.0673
3.1261 | -97.0732
3.1298 | -97.0931
3.1335 | -97.1184
3.1372 | -97.1459
3.1409 | -97.1643
3.1446 | -97.1606
3.1483 | -97.1245
3.152 | -97.0674
3.1557 | -97.0168
3.1594 | -96.9746
3.1631 | -96.9536
3.1668 | -96.9337
3.1705 | -96.9097
3.1742 | -96.8989
3.1779 | -96.9025
3.1816 | -96.9351
3.1853 | -96.9948
3.189 | -97.0617
3.1927 | -97.1116
3.1964 | -97.1224
3.2001 | -97.0781
3.2039 | -96.9384
3.2076 | -96.8367
3.2113 | -96.7346
3.215 | -96.6597
3.2187 | -96.6247
3.2224 | -96.6394
3.2261 | -96.6992
3.2298 | -96.7894
3.2335 | -96.8814
3.2372 | -96.9508
3.2409 | -96.9952
3.2446 | -97.0329
3.2483 | -97.0712
3.252 | -97.1212
3.2557 | -97.1882
3.2594 | -97.2626
3.2631 | -97.3333
3.2668 | -97.3776
3.2705 | -97.3905
3.2742 | -97.3834
3.2779 | -97.3561
3.2816 | -97.3171
3.2853 | -97.28
3.289 | -97.2527
3.2927 | -97.2448
3.2964 | -97.2541
3.3002 | -97.2785
3.3039 | -97.3152
3.3076 | -97.3438
3.3113 | -97.3531
3.315 | -97.3412
3.3187 | -97.3169
3.3224 | -97.2992
3.3261 | -97.2966
3.3298 | -97.325
3.3335 | -97.3891
3.3372 | -97.471
3.3409 | -97.5534
3.3446 | -97.6232
3.3483 | -97.6576
3.352 | -97.6751
3.3557 | -97.6787
3.3594 | -97.6815
3.3631 | -97.6858
3.3668 | -97.6908
3.3705 | -97.693
3.3742 | -97.6806
3.3779 | -97.6467
3.3816 | -97.5879
3.3853 | -97.5236
3.389 | -97.4779
3.3927 | -97.4711
3.3965 | -97.4963
3.4002 | -97.5203
3.4039 | -97.5355
3.4076 | -97.529
3.4113 | -97.499
3.415 | -97.4445
3.4187 | -97.3867
3.4224 | -97.328
3.4261 | -97.284
3.4298 | -97.2748
3.4335 | -97.2947
3.4372 | -97.3382
3.4409 | -97.3805
3.4446 | -97.4124
3.4483 | -97.4298
3.452 | -97.4434
3.4557 | -97.4628
3.4594 | -97.4844
3.4631 | -97.5023
3.4668 | -97.5256
3.4705 | -97.5601
3.4742 | -97.5961
3.4779 | -97.6497
3.4816 | -97.7206
3.4853 | -97.7909
3.4891 | -97.8744
3.4928 | -97.961
3.4965 | -98.0545
3.5002 | -98.152
3.5039 | -98.246
3.5076 | -98.3266
3.5113 | -98.3888
3.515 | -98.4222
3.5187 | -98.4283
3.5224 | -98.4065
3.5261 | -98.3655
3.5298 | -98.306
3.5335 | -98.2358
3.5372 | -98.1789
3.5409 | -98.154
3.5446 | -98.1455
3.5483 | -98.1507
3.552 | -98.1413
3.5557 | -98.1014
3.5594 | -98.0363
3.5631 | -97.9582
3.5668 | -97.8913
3.5705 | -97.8486
3.5742 | -97.8447
3.5779 | -97.8768
3.5816 | -97.9228
3.5854 | -97.9611
3.5891 | -97.9789
3.5928 | -97.971
3.5965 | -97.9489
3.6002 | -97.9471
3.6039 | -97.9782
3.6076 | -98.0586
3.6113 | -98.1823
3.615 | -98.3227
3.6187 | -98.4657
3.6224 | -98.5953
3.6261 | -98.7075
3.6298 | -98.7847
3.6335 | -98.8193
3.6372 | -98.8043
3.6409 | -98.7306
3.6446 | -98.5943
3.6483 | -98.4192
3.652 | -98.2299
3.6557 | -98.0449
3.6594 | -97.8829
3.6631 | -97.7586
3.6668 | -97.6842
3.6705 | -97.6704
3.6742 | -97.7182
3.6779 | -97.8228
3.6817 | -97.9598
3.6854 | -98.1035
3.6891 | -98.2406
3.6928 | -98.3516
3.6965 | -98.431
3.7002 | -98.4705
3.7039 | -98.473
3.7076 | -98.446
3.7113 | -98.3994
3.715 | -98.345
3.7187 | -98.2947
3.7224 | -98.2351
3.7261 | -98.1777
3.7298 | -98.1229
3.7335 | -98.062
3.7372 | -97.9806
3.7409 | -97.9093
3.7446 | -97.8441
3.7483 | -97.7859
3.752 | -97.7486
3.7557 | -97.7452
3.7594 | -97.7875
3.7631 | -97.8718
3.7668 | -97.9957
3.7705 | -98.1488
3.7742 | -98.3305
3.778 | -98.5146
3.7817 | -98.7007
3.7854 | -98.853
3.7891 | -98.9606
3.7928 | -99.0377
3.7965 | -99.0657
3.8002 | -99.0552
3.8039 | -99.0106
3.8076 | -98.9342
3.8113 | -98.8508
3.815 | -98.7831
3.8187 | -98.7372
3.8224 | -98.722
3.8261 | -98.7407
3.8298 | -98.7919
3.8335 | -98.8575
3.8372 | -98.9312
3.8409 | -99.0108
3.8446 | -99.0732
3.8483 | -99.1204
3.852 | -99.1455
3.8557 | -99.1464
3.8594 | -99.1407
3.8631 | -99.1268
3.8668 | -99.102
3.8706 | -99.0734
3.8743 | -99.0272
3.878 | -98.9654
3.8817 | -98.8964
3.8854 | -98.8146
3.8891 | -98.7378
3.8928 | -98.6656
3.8965 | -98.6071
3.9002 | -98.5667
3.9039 | -98.558
3.9076 | -98.5879
3.9113 | -98.641
3.915 | -98.7277
3.9187 | -98.84
3.9224 | -98.9736
3.9261 | -99.1204
3.9298 | -99.2695
3.9335 | -99.4123
3.9372 | -99.5198
3.9409 | -99.573
3.9446 | -99.5683
3.9483 | -99.5098
3.952 | -99.4008
3.9557 | -99.2658
3.9594 | -99.1253
3.9631 | -98.9832
3.9668 | -98.8606
3.9706 | -98.7782
3.9743 | -98.7506
3.978 | -98.7706
3.9817 | -98.8343
3.9854 | -98.9318
3.9891 | -99.0389
3.9928 | -99.1261
3.9965 | -99.1681
4.0002 | -99.1573
4.0039 | -99.1019
4.0076 | -99.0229
4.0113 | -98.9435
4.015 | -98.8819
4.0187 | -98.8411
4.0224 | -98.8275
4.0261 | -98.8365
4.0298 | -98.8434
4.0335 | -98.8467
4.0372 | -98.8413
4.0409 | -98.8175
4.0446 | -98.7754
4.0483 | -98.7261
4.052 | -98.6796
4.0557 | -98.6538
4.0594 | -98.6564
4.0632 | -98.701
4.0669 | -98.7741
4.0706 | -98.8729
4.0743 | -98.9931
4.078 | -99.1115
4.0817 | -99.2179
4.0854 | -99.2901
4.0891 | -99.3293
4.0928 | -99.3525
4.0965 | -99.3634
4.1002 | -99.3857
4.1039 | -99.4365
4.1076 | -99.5097
4.1113 | -99.5856
4.115 | -99.6678
4.1187 | -99.7365
4.1224 | -99.7878
4.1261 | -99.8062
4.1298 | -99.8067
4.1335 | -99.7878
4.1372 | -99.7554
4.1409 | -99.7315
4.1446 | -99.7139
4.1483 | -99.7015
4.152 | -99.6863
4.1557 | -99.6553
4.1595 | -99.6146
4.1632 | -99.5736
4.1669 | -99.5525
4.1706 | -99.566
4.1743 | -99.6224
4.178 | -99.7116
4.1817 | -99.819
4.1854 | -99.9095
4.1891 | -99.9651
4.1928 | -99.9535
4.1965 | -99.8823
4.2002 | -99.7678
4.2039 | -99.6185
4.2076 | -99.4676
4.2113 | -99.3295
4.215 | -99.2184
4.2187 | -99.152
4.2224 | -99.1322
4.2261 | -99.1643
4.2298 | -99.2383
4.2335 | -99.3446
4.2372 | -99.4576
4.2409 | -99.5587
4.2446 | -99.6338
4.2483 | -99.6889
4.252 | -99.7237
4.2558 | -99.7521
4.2595 | -99.7732
4.2632 | -99.7748
4.2669 | -99.7649
4.2706 | -99.7488
4.2743 | -99.7364
4.278 | -99.7231
4.2817 | -99.7119
4.2854 | -99.7063
4.2891 | -99.7066
4.2928 | -99.7075
4.2965 | -99.7063
4.3002 | -99.7041
4.3039 | -99.71
4.3076 | -99.7184
4.3113 | -99.7311
4.315 | -99.7378
4.3187 | -99.7318
4.3224 | -99.7193
4.3261 | -99.6978
4.3298 | -99.677
4.3335 | -99.655
4.3372 | -99.6264
4.3409 | -99.607
4.3446 | -99.5948
4.3483 | -99.5917
4.3521 | -99.5978
4.3558 | -99.5995
4.3595 | -99.5995
4.3632 | -99.5978
4.3669 | -99.5972
4.3706 | -99.5946
4.3743 | -99.5856
4.378 | -99.569
4.3817 | -99.5458
4.3854 | -99.5135
4.3891 | -99.4775
4.3928 | -99.437
4.3965 | -99.4021
4.4002 | -99.3808
4.4039 | -99.3753
4.4076 | -99.3892
4.4113 | -99.4154
4.415 | -99.4583
4.4187 | -99.5168
4.4224 | -99.587
4.4261 | -99.6683
4.4298 | -99.753
4.4335 | -99.8298
4.4372 | -99.9071
4.4409 | -99.9709
4.4447 | -100.0176
4.4484 | -100.0502
4.4521 | -100.0579
4.4558 | -100.046
4.4595 | -100.0216
4.4632 | -99.9877
4.4669 | -99.9471
4.4706 | -99.912
4.4743 | -99.8828
4.478 | -99.8745
4.4817 | -99.8919
4.4854 | -99.9207
4.4891 | -99.96
4.4928 | -100.0044
4.4965 | -100.0574
4.5002 | -100.1024
4.5039 | -100.1392
4.5076 | -100.1471
4.5113 | -100.139
4.515 | -100.1259
4.5187 | -100.1126
4.5224 | -100.1049
4.5261 | -100.1033
4.5298 | -100.1073
4.5335 | -100.1134
4.5372 | -100.1253
4.541 | -100.1351
4.5447 | -100.1388
4.5484 | -100.128
4.5521 | -100.1002
4.5558 | -100.0636
4.5595 | -100.0175
4.5632 | -99.9638
4.5669 | -99.9023
4.5706 | -99.8591
4.5743 | -99.8339
4.578 | -99.8289
4.5817 | -99.8427
4.5854 | -99.8573
4.5891 | -99.873
4.5928 | -99.8881
4.5965 | -99.9015
4.6002 | -99.9058
4.6039 | -99.9138
4.6076 | -99.9267
4.6113 | -99.9554
4.615 | -100.0176
4.6187 | -100.1061
4.6224 | -100.2201
4.6261 | -100.3496
4.6298 | -100.4684
4.6335 | -100.5642
4.6373 | -100.627
4.641 | -100.6386
4.6447 | -100.6197
4.6484 | -100.5895
4.6521 | -100.5456
4.6558 | -100.4881
4.6595 | -100.4179
4.6632 | -100.3406
4.6669 | -100.2538
4.6706 | -100.1586
4.6743 | -100.0551
4.678 | -99.957
4.6817 | -99.8727
4.6854 | -99.8108
4.6891 | -99.7805
4.6928 | -99.7788
4.6965 | -99.7973
4.7002 | -99.8282
4.7039 | -99.8693
4.7076 | -99.9067
4.7113 | -99.9384
4.715 | -99.9664
4.7187 | -99.9933
4.7224 | -100.0286
4.7261 | -100.0745
4.7298 | -100.1233
4.7336 | -100.1662
4.7373 | -100.2159
4.741 | -100.2701
4.7447 | -100.3216
4.7484 | -100.3711
4.7521 | -100.4146
4.7558 | -100.4564
4.7595 | -100.493
4.7632 | -100.5172
4.7669 | -100.5325
4.7706 | -100.5375
4.7743 | -100.5268
4.778 | -100.506
4.7817 | -100.4799
4.7854 | -100.4553
4.7891 | -100.4424
4.7928 | -100.4419
4.7965 | -100.4629
4.8002 | -100.5176
4.8039 | -100.5927
4.8076 | -100.6969
4.8113 | -100.7809
4.815 | -100.8721
4.8187 | -100.967
4.8224 | -101.0589
4.8261 | -101.1461
4.8299 | -101.2184
4.8336 | -101.2828
4.8373 | -101.3368
4.841 | -101.376
4.8447 | -101.3896
4.8484 | -101.3714
4.8521 | -101.3365
4.8558 | -101.2797
4.8595 | -101.2034
4.8632 | -101.1219
4.8669 | -101.0553
4.8706 | -101.0016
4.8743 | -100.9538
4.878 | -100.9267
4.8817 | -100.9242
4.8854 | -100.9374
4.8891 | -100.9654
4.8928 | -100.9838
4.8965 | -101.0023
4.9002 | -101.0017
4.9039 | -100.9799
4.9076 | -100.9461
4.9113 | -100.8954
4.915 | -100.8262
4.9187 | -100.7505
4.9225 | -100.6796
4.9262 | -100.6171
4.9299 | -100.5608
4.9336 | -100.5142
4.9373 | -100.4764
4.941 | -100.4551
4.9447 | -100.4348
4.9484 | -100.4171
4.9521 | -100.4108
4.9558 | -100.4114
4.9595 | -100.3997
4.9632 | -100.3858
4.9669 | -100.3772
4.9706 | -100.3718
4.9743 | -100.3599
4.978 | -100.3474
4.9817 | -100.3283
4.9854 | -100.308
4.9891 | -100.293
4.9928 | -100.2852
4.9965 | -100.2762
5.0002 | -100.2775
5.0039 | -100.2901
5.0076 | -100.3061
5.0113 | -100.3357
5.015 | -100.3675
5.0188 | -100.395
5.0225 | -100.4258
5.0262 | -100.4512
5.0299 | -100.4699
5.0336 | -100.4892
5.0373 | -100.5131
5.041 | -100.5402
5.0447 | -100.571
5.0484 | -100.6119
5.0521 | -100.6643
5.0558 | -100.7268
5.0595 | -100.7966
5.0632 | -100.8569
5.0669 | -100.9048
5.0706 | -100.9529
5.0743 | -100.992
5.078 | -101.0091
5.0817 | -101.0101
5.0854 | -101.0062
5.0891 | -100.9865
5.0928 | -100.9694
5.0965 | -100.9296
5.1002 | -100.8756
5.1039 | -100.8126
5.1076 | -100.7235
5.1113 | -100.6268
5.1151 | -100.5339
5.1188 | -100.4475
5.1225 | -100.3749
5.1262 | -100.3256
5.1299 | -100.3004
5.1336 | -100.3023
5.1373 | -100.3327
5.141 | -100.3844
5.1447 | -100.4634
5.1484 | -100.5541
5.1521 | -100.6496
5.1558 | -100.7539
5.1595 | -100.8463
5.1632 | -100.9148
5.1669 | -100.9532
5.1706 | -100.9542
5.1743 | -100.9321
5.178 | -100.8845
5.1817 | -100.8236
5.1854 | -100.7582
5.1891 | -100.6932
5.1928 | -100.6443
5.1965 | -100.6131
5.2002 | -100.6038
5.2039 | -100.6112
5.2076 | -100.6307
5.2114 | -100.6674
5.2151 | -100.7116
5.2188 | -100.7604
5.2225 | -100.8056
5.2262 | -100.8586
5.2299 | -100.9128
5.2336 | -100.9628
5.2373 | -101.0088
5.241 | -101.0421
5.2447 | -101.0622
5.2484 | -101.0717
5.2521 | -101.0705
5.2558 | -101.0533
5.2595 | -101.028
5.2632 | -101.0103
5.2669 | -100.9927
5.2706 | -100.9731
5.2743 | -100.9668
5.278 | -100.9698
5.2817 | -100.9866
5.2854 | -101.0192
5.2891 | -101.0649
5.2928 | -101.1275
5.2965 | -101.2133
5.3002 | -101.3148
5.3039 | -101.4135
5.3077 | -101.5179
5.3114 | -101.6108
5.3151 | -101.6881
5.3188 | -101.7491
5.3225 | -101.7882
5.3262 | -101.8048
5.3299 | -101.8049
5.3336 | -101.7901
5.3373 | -101.7697
5.341 | -101.737
5.3447 | -101.7092
5.3484 | -101.6822
5.3521 | -101.6566
5.3558 | -101.6236
5.3595 | -101.5915
5.3632 | -101.5589
5.3669 | -101.5206
5.3706 | -101.485
5.3743 | -101.4589
5.378 | -101.4381
5.3817 | -101.4301
5.3854 | -101.4348
5.3891 | -101.4482
5.3928 | -101.4727
5.3965 | -101.5086
5.4002 | -101.5393
5.404 | -101.5624
5.4077 | -101.5735
5.4114 | -101.5759
5.4151 | -101.5724
5.4188 | -101.5523
5.4225 | -101.5192
5.4262 | -101.4749
5.4299 | -101.4181
5.4336 | -101.3526
5.4373 | -101.2848
5.441 | -101.227
5.4447 | -101.1765
5.4484 | -101.1425
5.4521 | -101.1262
5.4558 | -101.1249
5.4595 | -101.1373
5.4632 | -101.1631
5.4669 | -101.2059
5.4706 | -101.2604
5.4743 | -101.3229
5.478 | -101.3909
5.4817 | -101.466
5.4854 | -101.5528
5.4891 | -101.6325
5.4928 | -101.7112
5.4966 | -101.7696
5.5003 | -101.8085
5.504 | -101.8307
5.5077 | -101.8376
5.5114 | -101.8238
5.5151 | -101.8014
5.5188 | -101.7654
5.5225 | -101.7288
5.5262 | -101.6881
5.5299 | -101.6471
5.5336 | -101.6124
5.5373 | -101.5871
5.541 | -101.5719
5.5447 | -101.5686
5.5484 | -101.571
5.5521 | -101.5709
5.5558 | -101.5755
5.5595 | -101.5747
5.5632 | -101.5716
5.5669 | -101.563
5.5706 | -101.5512
5.5743 | -101.5418
5.578 | -101.519
5.5817 | -101.4953
5.5854 | -101.4559
5.5891 | -101.4075
5.5929 | -101.3508
5.5966 | -101.2782
5.6003 | -101.19
5.604 | -101.1047
5.6077 | -101.0183
5.6114 | -100.9367
5.6151 | -100.8716
5.6188 | -100.8213
5.6225 | -100.7942
5.6262 | -100.7916
5.6299 | -100.8123
5.6336 | -100.8621
5.6373 | -100.9387
5.641 | -101.0316
5.6447 | -101.1345
5.6484 | -101.2401
5.6521 | -101.3437
5.6558 | -101.432
5.6595 | -101.4999
5.6632 | -101.5491
5.6669 | -101.573
5.6706 | -101.5743
5.6743 | -101.5514
5.678 | -101.5165
5.6817 | -101.4765
5.6854 | -101.4364
5.6892 | -101.389
5.6929 | -101.3488
5.6966 | -101.3149
5.7003 | -101.2827
5.704 | -101.2648
5.7077 | -101.256
5.7114 | -101.2642
5.7151 | -101.2859
5.7188 | -101.3138
5.7225 | -101.3546
5.7262 | -101.4099
5.7299 | -101.4684
5.7336 | -101.5221
5.7373 | -101.5695
5.741 | -101.6194
5.7447 | -101.6612
5.7484 | -101.6863
5.7521 | -101.7013
5.7558 | -101.7027
5.7595 | -101.6914
5.7632 | -101.6563
5.7669 | -101.6092
5.7706 | -101.551
5.7743 | -101.4902
5.778 | -101.431
5.7817 | -101.385
5.7855 | -101.3527
5.7892 | -101.3361
5.7929 | -101.3371
5.7966 | -101.3468
5.8003 | -101.3672
5.804 | -101.3954
5.8077 | -101.417
5.8114 | -101.4375
5.8151 | -101.4568
5.8188 | -101.4728
5.8225 | -101.4788
5.8262 | -101.4792
5.8299 | -101.4801
5.8336 | -101.4826
5.8373 | -101.4842
5.841 | -101.4935
5.8447 | -101.5021
5.8484 | -101.5139
5.8521 | -101.5211
5.8558 | -101.5276
5.8595 | -101.5378
5.8632 | -101.536
5.8669 | -101.5295
5.8706 | -101.5244
5.8743 | -101.5178
5.878 | -101.508
5.8818 | -101.4975
5.8855 | -101.4899
5.8892 | -101.4855
5.8929 | -101.4888
5.8966 | -101.4991
5.9003 | -101.5098
5.904 | -101.5258
5.9077 | -101.5454
5.9114 | -101.5653
5.9151 | -101.5832
5.9188 | -101.5915
5.9225 | -101.5882
5.9262 | -101.5651
5.9299 | -101.5215
5.9336 | -101.4627
5.9373 | -101.3857
5.941 | -101.2957
5.9447 | -101.1902
5.9484 | -101.0789
5.9521 | -100.9825
5.9558 | -100.8862
5.9595 | -100.7949
5.9632 | -100.7174
5.9669 | -100.6623
5.9706 | -100.6288
5.9744 | -100.6133
5.9781 | -100.6179
5.9818 | -100.6422
5.9855 | -100.6756
5.9892 | -100.7213
5.9929 | -100.7812
5.9966 | -100.8454
6.0003 | -100.9093
6.004 | -100.9726
6.0077 | -101.0279
6.0114 | -101.0867
6.0151 | -101.1451
6.0188 | -101.2072
6.0225 | -101.27
6.0262 | -101.3363
6.0299 | -101.4057
6.0336 | -101.473
6.0373 | -101.5418
6.041 | -101.6073
6.0447 | -101.6632
6.0484 | -101.7213
6.0521 | -101.7711
6.0558 | -101.8048
6.0595 | -101.8327
6.0632 | -101.8594
6.0669 | -101.885
6.0707 | -101.8957
6.0744 | -101.9009
6.0781 | -101.9071
6.0818 | -101.9147
6.0855 | -101.9165
6.0892 | -101.919
6.0929 | -101.9242
6.0966 | -101.9315
6.1003 | -101.9408
6.104 | -101.9504
6.1077 | -101.9621
6.1114 | -101.9723
6.1151 | -101.9848
6.1188 | -101.9946
6.1225 | -102.0056
6.1262 | -102.0184
6.1299 | -102.0319
6.1336 | -102.0417
6.1373 | -102.0527
6.141 | -102.0625
6.1447 | -102.0691
6.1484 | -102.0757
6.1521 | -102.0734
6.1558 | -102.0595
6.1595 | -102.0332
6.1632 | -102.0086
6.167 | -101.9779
6.1707 | -101.9387
6.1744 | -101.889
6.1781 | -101.8423
6.1818 | -101.8038
6.1855 | -101.7713
6.1892 | -101.7508
6.1929 | -101.7426
6.1966 | -101.7499
6.2003 | -101.7657
6.204 | -101.7894
6.2077 | -101.8185
6.2114 | -101.849
6.2151 | -101.8704
6.2188 | -101.8888
6.2225 | -101.9064
6.2262 | -101.9222
6.2299 | -101.9371
6.2336 | -101.9582
6.2373 | -101.9836
6.241 | -102.0078
6.2447 | -102.0368
6.2484 | -102.0763
6.2521 | -102.1197
6.2558 | -102.1564
6.2595 | -102.198
6.2633 | -102.2342
6.267 | -102.2644
6.2707 | -102.2855
6.2744 | -102.2905
6.2781 | -102.2847
6.2818 | -102.2693
6.2855 | -102.2448
6.2892 | -102.2151
6.2929 | -102.1923
6.2966 | -102.1768
6.3003 | -102.1751
6.304 | -102.1896
6.3077 | -102.2268
6.3114 | -102.2728
6.3151 | -102.3327
6.3188 | -102.4056
6.3225 | -102.4755
6.3262 | -102.5361
6.3299 | -102.5862
6.3336 | -102.626
6.3373 | -102.6515
6.341 | -102.6575
6.3447 | -102.6555
6.3484 | -102.6397
6.3521 | -102.6061
6.3559 | -102.5733
6.3596 | -102.5301
6.3633 | -102.4795
6.367 | -102.4289
6.3707 | -102.3787
6.3744 | -102.324
6.3781 | -102.2742
6.3818 | -102.229
6.3855 | -102.188
6.3892 | -102.1584
6.3929 | -102.1354
6.3966 | -102.1249
6.4003 | -102.1225
6.404 | -102.1272
6.4077 | -102.1431
6.4114 | -102.1784
6.4151 | -102.2038
6.4188 | -102.2344
6.4225 | -102.2619
6.4262 | -102.281
6.4299 | -102.2921
6.4336 | -102.2929
6.4373 | -102.2924
6.441 | -102.2903
6.4447 | -102.2854
6.4484 | -102.2776
6.4522 | -102.2727
6.4559 | -102.2755
6.4596 | -102.2786
6.4633 | -102.2834
6.467 | -102.2969
6.4707 | -102.3217
6.4744 | -102.3467
6.4781 | -102.3738
6.4818 | -102.4042
6.4855 | -102.4381
6.4892 | -102.4771
6.4929 | -102.5139
6.4966 | -102.5564
6.5003 | -102.5998
6.504 | -102.6306
6.5077 | -102.6522
6.5114 | -102.6625
6.5151 | -102.6628
6.5188 | -102.653
6.5225 | -102.6386
6.5262 | -102.622
6.5299 | -102.5937
6.5336 | -102.5605
6.5373 | -102.5243
6.541 | -102.4897
6.5447 | -102.464
6.5485 | -102.4388
6.5522 | -102.413
6.5559 | -102.3916
6.5596 | -102.3809
6.5633 | -102.3696
6.567 | -102.3557
6.5707 | -102.338
6.5744 | -102.3182
6.5781 | -102.2946
6.5818 | -102.2706
6.5855 | -102.2448
6.5892 | -102.2216
6.5929 | -102.205
6.5966 | -102.1846
6.6003 | -102.1698
6.604 | -102.1696
6.6077 | -102.1788
6.6114 | -102.2034
6.6151 | -102.2345
6.6188 | -102.27
6.6225 | -102.3092
6.6262 | -102.3533
6.6299 | -102.4015
6.6336 | -102.449
6.6373 | -102.498
6.641 | -102.5474
6.6448 | -102.5946
6.6485 | -102.646
6.6522 | -102.6982
6.6559 | -102.7499
6.6596 | -102.7989
6.6633 | -102.8531
6.667 | -102.9072
6.6707 | -102.9619
6.6744 | -103.0172
6.6781 | -103.067
6.6818 | -103.1071
6.6855 | -103.1277
6.6892 | -103.1409
6.6929 | -103.1428
6.6966 | -103.1428
6.7003 | -103.1358
6.704 | -103.1132
6.7077 | -103.0883
6.7114 | -103.0605
6.7151 | -103.0342
6.7188 | -103.0151
6.7225 | -102.9985
6.7262 | -102.9875
6.7299 | -102.9896
6.7336 | -102.9973
6.7373 | -103.0096
6.7411 | -103.0288
6.7448 | -103.0541
6.7485 | -103.0793
6.7522 | -103.1032
6.7559 | -103.1245
6.7596 | -103.1358
6.7633 | -103.138
6.767 | -103.1287
6.7707 | -103.108
6.7744 | -103.0855
6.7781 | -103.0567
6.7818 | -103.0215
6.7855 | -102.989
6.7892 | -102.9488
6.7929 | -102.9106
6.7966 | -102.8644
6.8003 | -102.8143
6.804 | -102.7699
6.8077 | -102.7261
6.8114 | -102.6858
6.8151 | -102.6501
6.8188 | -102.6185
6.8225 | -102.5908
6.8262 | -102.5701
6.8299 | -102.5567
6.8336 | -102.5519
6.8374 | -102.5542
6.8411 | -102.5631
6.8448 | -102.5793
6.8485 | -102.6057
6.8522 | -102.6444
6.8559 | -102.6848
6.8596 | -102.7334
6.8633 | -102.7838
6.867 | -102.8317
6.8707 | -102.8733
6.8744 | -102.9051
6.8781 | -102.9325
6.8818 | -102.9526
6.8855 | -102.963
6.8892 | -102.9649
6.8929 | -102.9582
6.8966 | -102.9537
6.9003 | -102.9509
6.904 | -102.9386
6.9077 | -102.9248
6.9114 | -102.9069
6.9151 | -102.8911
6.9188 | -102.8783
6.9225 | -102.8663
6.9262 | -102.8505
6.9299 | -102.8344
6.9337 | -102.8258
6.9374 | -102.8225
6.9411 | -102.8236
6.9448 | -102.8284
6.9485 | -102.8305
6.9522 | -102.8337
6.9559 | -102.8338
6.9596 | -102.8449
6.9633 | -102.8576
6.967 | -102.8725
6.9707 | -102.8891
6.9744 | -102.9079
6.9781 | -102.9263
6.9818 | -102.952
6.9855 | -102.9806
6.9892 | -103.0108
6.9929 | -103.046
6.9966 | -103.086
7.0003 | -103.1256
7.004 | -103.1651
7.0077 | -103.2121
7.0114 | -103.2612
7.0151 | -103.3107
7.0188 | -103.3566
7.0225 | -103.3962
7.0263 | -103.4308
7.03 | -103.4547
7.0337 | -103.472
7.0374 | -103.4856
7.0411 | -103.493
7.0448 | -103.4941
7.0485 | -103.4915
7.0522 | -103.4832
7.0559 | -103.4648
7.0596 | -103.4389
7.0633 | -103.4035
7.067 | -103.3544
7.0707 | -103.2982
7.0744 | -103.2361
7.0781 | -103.1741
7.0818 | -103.1061
7.0855 | -103.0333
7.0892 | -102.9595
7.0929 | -102.8959
7.0966 | -102.8507
7.1003 | -102.8147
7.104 | -102.7915
7.1077 | -102.7827
7.1114 | -102.7902
7.1151 | -102.8098
7.1188 | -102.8354
7.1226 | -102.8739
7.1263 | -102.9192
7.13 | -102.9693
7.1337 | -103.0191
7.1374 | -103.0678
7.1411 | -103.1094
7.1448 | -103.1393
7.1485 | -103.1606
7.1522 | -103.1803
7.1559 | -103.1934
7.1596 | -103.2026
7.1633 | -103.2103
7.167 | -103.208
7.1707 | -103.1993
7.1744 | -103.1874
7.1781 | -103.175
7.1818 | -103.1576
7.1855 | -103.1392
7.1892 | -103.1259
7.1929 | -103.1091
7.1966 | -103.0891
7.2003 | -103.0645
7.204 | -103.029
7.2077 | -102.9914
7.2114 | -102.9458
7.2151 | -102.9031
7.2189 | -102.8633
7.2226 | -102.8245
7.2263 | -102.7954
7.23 | -102.7734
7.2337 | -102.762
7.2374 | -102.7608
7.2411 | -102.7725
7.2448 | -102.7933
7.2485 | -102.8211
7.2522 | -102.8606
7.2559 | -102.9034
7.2596 | -102.9432
7.2633 | -102.9908
7.267 | -103.0349
7.2707 | -103.0781
7.2744 | -103.1236
7.2781 | -103.1637
7.2818 | -103.1943
7.2855 | -103.2154
7.2892 | -103.2318
7.2929 | -103.2406
7.2966 | -103.2419
7.3003 | -103.2488
7.304 | -103.2612
7.3077 | -103.2762
7.3114 | -103.2992
7.3152 | -103.3361
7.3189 | -103.3807
7.3226 | -103.4295
7.3263 | -103.4857
7.33 | -103.5516
7.3337 | -103.6151
7.3374 | -103.6671
7.3411 | -103.7047
7.3448 | -103.7323
7.3485 | -103.7527
7.3522 | -103.7642
7.3559 | -103.7588
7.3596 | -103.739
7.3633 | -103.7173
7.367 | -103.6992
7.3707 | -103.6709
7.3744 | -103.6399
7.3781 | -103.6075
7.3818 | -103.5761
7.3855 | -103.5474
7.3892 | -103.5182
7.3929 | -103.4821
7.3966 | -103.4445
7.4003 | -103.4038
7.404 | -103.3653
7.4078 | -103.3232
7.4115 | -103.2843
7.4152 | -103.2512
7.4189 | -103.2114
7.4226 | -103.1684
7.4263 | -103.1354
7.43 | -103.1091
7.4337 | -103.0839
7.4374 | -103.0586
7.4411 | -103.0304
7.4448 | -103.007
7.4485 | -102.993
7.4522 | -102.9758
7.4559 | -102.9556
7.4596 | -102.9405
7.4633 | -102.9325
7.467 | -102.9349
7.4707 | -102.9475
7.4744 | -102.9623
7.4781 | -103.0054
7.4818 | -103.0396
7.4855 | -103.0812
7.4892 | -103.1229
7.4929 | -103.1639
7.4966 | -103.2012
7.5003 | -103.2381
7.5041 | -103.2741
7.5078 | -103.3033
7.5115 | -103.331
7.5152 | -103.3582
7.5189 | -103.3866
7.5226 | -103.4081
7.5263 | -103.4275
7.53 | -103.4535
7.5337 | -103.4771
7.5374 | -103.4898
7.5411 | -103.4952
7.5448 | -103.5024
7.5485 | -103.5137
7.5522 | -103.5289
7.5559 | -103.5371
7.5596 | -103.5369
7.5633 | -103.5343
7.567 | -103.5257
7.5707 | -103.5097
7.5744 | -103.4976
7.5781 | -103.4791
7.5818 | -103.4538
7.5855 | -103.434
7.5892 | -103.4204
7.5929 | -103.4088
7.5966 | -103.403
7.6004 | -103.4063
7.6041 | -103.4167
7.6078 | -103.4363
7.6115 | -103.4607
7.6152 | -103.4892
7.6189 | -103.5233
7.6226 | -103.5638
7.6263 | -103.6011
7.63 | -103.6352
7.6337 | -103.6681
7.6374 | -103.7068
7.6411 | -103.7462
7.6448 | -103.7836
7.6485 | -103.8163
7.6522 | -103.8445
7.6559 | -103.8641
7.6596 | -103.8721
7.6633 | -103.8668
7.667 | -103.8526
7.6707 | -103.8359
7.6744 | -103.8134
7.6781 | -103.7833
7.6818 | -103.7497
7.6855 | -103.7128
7.6892 | -103.6691
7.6929 | -103.6284
7.6967 | -103.5893
7.7004 | -103.5498
7.7041 | -103.5074
7.7078 | -103.4607
7.7115 | -103.4169
7.7152 | -103.3757
7.7189 | -103.3386
7.7226 | -103.3074
7.7263 | -103.2856
7.73 | -103.2715
7.7337 | -103.267
7.7374 | -103.2748
7.7411 | -103.2946
7.7448 | -103.3247
7.7485 | -103.3624
7.7522 | -103.4014
7.7559 | -103.4472
7.7596 | -103.4932
7.7633 | -103.5349
7.767 | -103.5727
7.7707 | -103.609
7.7744 | -103.6405
7.7781 | -103.6713
7.7818 | -103.692
7.7855 | -103.706
7.7893 | -103.7129
7.793 | -103.7174
7.7967 | -103.7179
7.8004 | -103.7122
7.8041 | -103.7034
7.8078 | -103.6848
7.8115 | -103.6663
7.8152 | -103.6524
7.8189 | -103.644
7.8226 | -103.6427
7.8263 | -103.6409
7.83 | -103.6397
7.8337 | -103.6411
7.8374 | -103.6421
7.8411 | -103.65
7.8448 | -103.6561
7.8485 | -103.6575
7.8522 | -103.6537
7.8559 | -103.6495
7.8596 | -103.6444
7.8633 | -103.634
7.867 | -103.6254
7.8707 | -103.6153
7.8744 | -103.612
7.8781 | -103.6178
7.8818 | -103.6291
7.8856 | -103.6459
7.8893 | -103.6693
7.893 | -103.6964
7.8967 | -103.7366
7.9004 | -103.7812
7.9041 | -103.8325
7.9078 | -103.8884
7.9115 | -103.9499
7.9152 | -104.015
7.9189 | -104.074
7.9226 | -104.1297
7.9263 | -104.1799
7.93 | -104.216
7.9337 | -104.2403
7.9374 | -104.2576
7.9411 | -104.2644
7.9448 | -104.2603
7.9485 | -104.2421
7.9522 | -104.2153
7.9559 | -104.1806
7.9596 | -104.1401
7.9633 | -104.1021
7.967 | -104.0685
7.9707 | -104.0386
7.9744 | -104.0045
7.9781 | -103.9798
7.9819 | -103.9625
7.9856 | -103.9442
7.9893 | -103.9208
7.993 | -103.8982
7.9967 | -103.8756
8.0004 | -103.8488
8.0041 | -103.8226
8.0078 | -103.7944
8.0115 | -103.7656
8.0152 | -103.7444
8.0189 | -103.722
8.0226 | -103.7054
8.0263 | -103.6964
8.03 | -103.6905
8.0337 | -103.6911
8.0374 | -103.7002
8.0411 | -103.7137
8.0448 | -103.726
8.0485 | -103.7344
8.0522 | -103.7429
8.0559 | -103.7497
8.0596 | -103.7549
8.0633 | -103.7598
8.067 | -103.7636
8.0707 | -103.7617
8.0744 | -103.7616
8.0782 | -103.7669
8.0819 | -103.7733
8.0856 | -103.7836
8.0893 | -103.7983
8.093 | -103.8215
8.0967 | -103.8451
8.1004 | -103.8713
8.1041 | -103.9043
8.1078 | -103.9414
8.1115 | -103.9791
8.1152 | -104.0205
8.1189 | -104.0591
8.1226 | -104.0979
8.1263 | -104.1363
8.13 | -104.1681
8.1337 | -104.198
8.1374 | -104.2226
8.1411 | -104.2384
8.1448 | -104.2419
8.1485 | -104.2351
8.1522 | -104.2273
8.1559 | -104.2161
8.1596 | -104.204
8.1633 | -104.1881
8.167 | -104.1689
8.1707 | -104.1493
8.1745 | -104.1316
8.1782 | -104.1085
8.1819 | -104.0928
8.1856 | -104.0747
8.1893 | -104.0536
8.193 | -104.0376
8.1967 | -104.0207
8.2004 | -104.009
8.2041 | -103.9965
8.2078 | -103.9826
8.2115 | -103.966
8.2152 | -103.9551
8.2189 | -103.9462
8.2226 | -103.9354
8.2263 | -103.9283
8.23 | -103.9214
8.2337 | -103.9151
8.2374 | -103.9064
8.2411 | -103.9014
8.2448 | -103.8958
8.2485 | -103.8904
8.2522 | -103.8901
8.2559 | -103.8901
8.2596 | -103.891
8.2633 | -103.8909
8.267 | -103.8948
8.2708 | -103.9
8.2745 | -103.9032
8.2782 | -103.9007
8.2819 | -103.8926
8.2856 | -103.887
8.2893 | -103.8813
8.293 | -103.8801
8.2967 | -103.8722
8.3004 | -103.8688
8.3041 | -103.8616
8.3078 | -103.8541
8.3115 | -103.8522
8.3152 | -103.8524
8.3189 | -103.8499
8.3226 | -103.8431
8.3263 | -103.8377
8.33 | -103.8365
8.3337 | -103.8359
8.3374 | -103.8354
8.3411 | -103.8317
8.3448 | -103.8343
8.3485 | -103.8377
8.3522 | -103.8457
8.3559 | -103.8498
8.3596 | -103.8568
8.3634 | -103.869
8.3671 | -103.8869
8.3708 | -103.9018
8.3745 | -103.915
8.3782 | -103.9334
8.3819 | -103.9508
8.3856 | -103.9633
8.3893 | -103.9753
8.393 | -103.9911
8.3967 | -104.0044
8.4004 | -104.0192
8.4041 | -104.0386
8.4078 | -104.0552
8.4115 | -104.0618
8.4152 | -104.061
8.4189 | -104.0626
8.4226 | -104.0621
8.4263 | -104.0609
8.43 | -104.0592
8.4337 | -104.0515
8.4374 | -104.0322
8.4411 | -104.0133
8.4448 | -103.9978
8.4485 | -103.9869
8.4522 | -103.9705
8.4559 | -103.9582
8.4597 | -103.9463
8.4634 | -103.94
8.4671 | -103.9343
8.4708 | -103.9356
8.4745 | -103.944
8.4782 | -103.9593
8.4819 | -103.9744
8.4856 | -103.9888
8.4893 | -103.9981
8.493 | -104.0181
8.4967 | -104.0409
8.5004 | -104.0582
8.5041 | -104.0742
8.5078 | -104.0949
8.5115 | -104.1105
8.5152 | -104.1137
8.5189 | -104.1165
8.5226 | -104.1174
8.5263 | -104.1062
8.53 | -104.0944
8.5337 | -104.085
8.5374 | -104.0642
8.5411 | -104.0386
8.5448 | -104.0104
8.5485 | -103.9816
8.5522 | -103.9531
8.556 | -103.9285
8.5597 | -103.907
8.5634 | -103.8904
8.5671 | -103.8791
8.5708 | -103.8672
8.5745 | -103.8553
8.5782 | -103.8487
8.5819 | -103.8452
8.5856 | -103.8477
8.5893 | -103.855
8.593 | -103.8675
8.5967 | -103.8813
8.6004 | -103.8954
8.6041 | -103.9164
8.6078 | -103.9471
8.6115 | -103.9789
8.6152 | -104.0134
8.6189 | -104.0468
8.6226 | -104.0805
8.6263 | -104.1166
8.63 | -104.1518
8.6337 | -104.184
8.6374 | -104.2108
8.6411 | -104.2373
8.6448 | -104.2596
8.6485 | -104.2764
8.6523 | -104.2937
8.656 | -104.3023
8.6597 | -104.305
8.6634 | -104.3082
8.6671 | -104.3054
8.6708 | -104.3007
8.6745 | -104.2999
8.6782 | -104.299
8.6819 | -104.2975
8.6856 | -104.2976
8.6893 | -104.3005
8.693 | -104.3063
8.6967 | -104.3141
8.7004 | -104.3245
8.7041 | -104.3397
8.7078 | -104.3556
8.7115 | -104.3715
8.7152 | -104.3808
8.7189 | -104.3879
8.7226 | -104.3899
8.7263 | -104.3941
8.73 | -104.4028
8.7337 | -104.4045
8.7374 | -104.3948
8.7411 | -104.3846
8.7448 | -104.3776
8.7486 | -104.36
8.7523 | -104.3397
8.756 | -104.318
8.7597 | -104.2968
8.7634 | -104.2714
8.7671 | -104.2471
8.7708 | -104.2188
8.7745 | -104.1902
8.7782 | -104.1638
8.7819 | -104.1366
8.7856 | -104.1147
8.7893 | -104.0975
8.793 | -104.0835
8.7967 | -104.0757
8.8004 | -104.0681
8.8041 | -104.061
8.8078 | -104.0606
8.8115 | -104.0645
8.8152 | -104.0712
8.8189 | -104.0807
8.8226 | -104.095
8.8263 | -104.1074
8.83 | -104.1223
8.8337 | -104.1365
8.8374 | -104.1497
8.8412 | -104.1635
8.8449 | -104.1776
8.8486 | -104.1897
8.8523 | -104.1978
8.856 | -104.1992
8.8597 | -104.1898
8.8634 | -104.1804
8.8671 | -104.1724
8.8708 | -104.1634
8.8745 | -104.1476
8.8782 | -104.1353
8.8819 | -104.1216
8.8856 | -104.1038
8.8893 | -104.0872
8.893 | -104.0716
8.8967 | -104.0583
8.9004 | -104.052
8.9041 | -104.0435
8.9078 | -104.0334
8.9115 | -104.0274
8.9152 | -104.0284
8.9189 | -104.0334
8.9226 | -104.0413
8.9263 | -104.0476
8.93 | -104.0613
8.9337 | -104.0758
8.9375 | -104.0873
8.9412 | -104.0989
8.9449 | -104.1047
8.9486 | -104.1079
8.9523 | -104.1099
8.956 | -104.112
8.9597 | -104.1079
8.9634 | -104.0999
8.9671 | -104.0913
8.9708 | -104.0766
8.9745 | -104.0631
8.9782 | -104.0491
8.9819 | -104.0383
8.9856 | -104.0233
8.9893 | -104.0103
8.993 | -103.9965
8.9967 | -103.9786
9.0004 | -103.9607
9.0041 | -103.9467
9.0078 | -103.9285
9.0115 | -103.9094
9.0152 | -103.8969
9.0189 | -103.8855
9.0226 | -103.8769
9.0263 | -103.8689
9.03 | -103.8603
9.0338 | -103.857
9.0375 | -103.8581
9.0412 | -103.8582
9.0449 | -103.8553
9.0486 | -103.8563
9.0523 | -103.8628
9.056 | -103.8689
9.0597 | -103.8721
9.0634 | -103.88
9.0671 | -103.893
9.0708 | -103.9137
9.0745 | -103.9406
9.0782 | -103.9671
9.0819 | -103.9948
9.0856 | -104.0174
9.0893 | -104.0395
9.093 | -104.0675
9.0967 | -104.0953
9.1004 | -104.1192
9.1041 | -104.1404
9.1078 | -104.1576
9.1115 | -104.1763
9.1152 | -104.1891
9.1189 | -104.2039
9.1226 | -104.2187
9.1263 | -104.2329
9.1301 | -104.248
9.1338 | -104.26
9.1375 | -104.2726
9.1412 | -104.2907
9.1449 | -104.3086
9.1486 | -104.3264
9.1523 | -104.3477
9.156 | -104.3733
9.1597 | -104.3979
9.1634 | -104.4241
9.1671 | -104.4563
9.1708 | -104.4853
9.1745 | -104.5173
9.1782 | -104.5541
9.1819 | -104.591
9.1856 | -104.6219
9.1893 | -104.6563
9.193 | -104.6891
9.1967 | -104.7185
9.2004 | -104.7393
9.2041 | -104.7533
9.2078 | -104.7637
9.2115 | -104.7714
9.2152 | -104.7788
9.2189 | -104.7871
9.2226 | -104.786
9.2264 | -104.7799
9.2301 | -104.7698
9.2338 | -104.7604
9.2375 | -104.755
9.2412 | -104.7445
9.2449 | -104.7319
9.2486 | -104.7154
9.2523 | -104.6978
9.256 | -104.6804
9.2597 | -104.6627
9.2634 | -104.645
9.2671 | -104.6296
9.2708 | -104.6157
9.2745 | -104.5988
9.2782 | -104.5838
9.2819 | -104.5716
9.2856 | -104.5602
9.2893 | -104.5506
9.293 | -104.5393
9.2967 | -104.5362
9.3004 | -104.5366
9.3041 | -104.534
9.3078 | -104.5337
9.3115 | -104.5313
9.3152 | -104.5304
9.3189 | -104.5319
9.3227 | -104.5394
9.3264 | -104.5461
9.3301 | -104.5519
9.3338 | -104.551
9.3375 | -104.5499
9.3412 | -104.5526
9.3449 | -104.5569
9.3486 | -104.562
9.3523 | -104.5647
9.356 | -104.5679
9.3597 | -104.5778
9.3634 | -104.591
9.3671 | -104.6027
9.3708 | -104.6149
9.3745 | -104.6265
9.3782 | -104.6433
9.3819 | -104.6602
9.3856 | -104.6771
9.3893 | -104.7006
9.393 | -104.7211
9.3967 | -104.7383
9.4004 | -104.7524
9.4041 | -104.7645
9.4078 | -104.7798
9.4115 | -104.7969
9.4153 | -104.8091
9.419 | -104.8151
9.4227 | -104.8197
9.4264 | -104.8185
9.4301 | -104.818
9.4338 | -104.8155
9.4375 | -104.8126
9.4412 | -104.8044
9.4449 | -104.7963
9.4486 | -104.7844
9.4523 | -104.7674
9.456 | -104.7524
9.4597 | -104.7372
9.4634 | -104.7194
9.4671 | -104.7013
9.4708 | -104.6933
9.4745 | -104.6862
9.4782 | -104.6741
9.4819 | -104.6639
9.4856 | -104.6592
9.4893 | -104.6564
9.493 | -104.6565
9.4967 | -104.6549
9.5004 | -104.6544
9.5041 | -104.6522
9.5078 | -104.6474
9.5116 | -104.6428
9.5153 | -104.6419
9.519 | -104.6372
9.5227 | -104.6272
9.5264 | -104.611
9.5301 | -104.5974
9.5338 | -104.5809
9.5375 | -104.563
9.5412 | -104.5439
9.5449 | -104.5209
9.5486 | -104.5004
9.5523 | -104.4836
9.556 | -104.4698
9.5597 | -104.4587
9.5634 | -104.4505
9.5671 | -104.4439
9.5708 | -104.4375
9.5745 | -104.4308
9.5782 | -104.431
9.5819 | -104.4358
9.5856 | -104.4421
9.5893 | -104.4477
9.593 | -104.4528
9.5967 | -104.4584
9.6004 | -104.4679
9.6041 | -104.4782
9.6079 | -104.4859
9.6116 | -104.4979
9.6153 | -104.5124
9.619 | -104.5237
9.6227 | -104.5392
9.6264 | -104.5562
9.6301 | -104.5738
9.6338 | -104.5894
9.6375 | -104.6075
9.6412 | -104.6265
9.6449 | -104.6431
9.6486 | -104.6559
9.6523 | -104.6635
9.656 | -104.667
9.6597 | -104.6694
9.6634 | -104.6672
9.6671 | -104.6636
9.6708 | -104.6569
9.6745 | -104.6506
9.6782 | -104.6405
9.6819 | -104.6252
9.6856 | -104.6113
9.6893 | -104.5987
9.693 | -104.5855
9.6967 | -104.5757
9.7004 | -104.5652
9.7042 | -104.5596
9.7079 | -104.5579
9.7116 | -104.5561
9.7153 | -104.5594
9.719 | -104.5642
9.7227 | -104.5739
9.7264 | -104.592
9.7301 | -104.6108
9.7338 | -104.6307
9.7375 | -104.6535
9.7412 | -104.677
9.7449 | -104.6993
9.7486 | -104.716
9.7523 | -104.7278
9.756 | -104.7335
9.7597 | -104.7407
9.7634 | -104.7451
9.7671 | -104.7417
9.7708 | -104.7322
9.7745 | -104.7204
9.7782 | -104.7085
9.7819 | -104.696
9.7856 | -104.6799
9.7893 | -104.6622
9.793 | -104.6466
9.7967 | -104.6346
9.8005 | -104.6229
9.8042 | -104.6135
9.8079 | -104.6069
9.8116 | -104.6042
9.8153 | -104.5995
9.819 | -104.5959
9.8227 | -104.5975
9.8264 | -104.6009
9.8301 | -104.6024
9.8338 | -104.6044
9.8375 | -104.6075
9.8412 | -104.6062
9.8449 | -104.6056
9.8486 | -104.6082
9.8523 | -104.6112
9.856 | -104.6138
9.8597 | -104.6118
9.8634 | -104.6117
9.8671 | -104.6088
9.8708 | -104.6043
9.8745 | -104.6029
9.8782 | -104.5981
9.8819 | -104.5899
9.8856 | -104.5806
9.8893 | -104.5675
9.8931 | -104.5573
9.8968 | -104.5474
9.9005 | -104.5392
9.9042 | -104.5285
9.9079 | -104.5128
9.9116 | -104.4977
9.9153 | -104.4811
9.919 | -104.4675
9.9227 | -104.4575
9.9264 | -104.4508
9.9301 | -104.4506
9.9338 | -104.4566
9.9375 | -104.4627
9.9412 | -104.4691
9.9449 | -104.4765
9.9486 | -104.4858
9.9523 | -104.499
9.956 | -104.5077
9.9597 | -104.5223
9.9634 | -104.5445
9.9671 | -104.5658
9.9708 | -104.5878
9.9745 | -104.6104
9.9782 | -104.6316
9.9819 | -104.6547
9.9856 | -104.6751
9.9894 | -104.6934
9.9931 | -104.7066
9.9968 | -104.7194
10.0005 | -104.7335
10.0042 | -104.7476
10.0079 | -104.7537
10.0116 | -104.7545
10.0153 | -104.7511
10.019 | -104.7464
10.0227 | -104.7396
10.0264 | -104.7254
10.0301 | -104.7102
10.0338 | -104.6975
10.0375 | -104.6822
10.0412 | -104.6659
10.0449 | -104.6493
10.0486 | -104.6313
10.0523 | -104.616
10.056 | -104.6006
10.0597 | -104.5913
10.0634 | -104.5805
10.0671 | -104.5657
10.0708 | -104.5531
10.0745 | -104.5437
10.0782 | -104.5403
10.0819 | -104.5402
10.0857 | -104.5426
10.0894 | -104.5496
10.0931 | -104.5624
10.0968 | -104.5789
10.1005 | -104.5988
10.1042 | -104.6183
10.1079 | -104.6331
10.1116 | -104.6535
10.1153 | -104.6768
10.119 | -104.7002
10.1227 | -104.7274
10.1264 | -104.7547
10.1301 | -104.783
10.1338 | -104.8095
10.1375 | -104.8342
10.1412 | -104.8542
10.1449 | -104.868
10.1486 | -104.8818
10.1523 | -104.8976
10.156 | -104.9092
10.1597 | -104.9209
10.1634 | -104.9335
10.1671 | -104.9469
10.1708 | -104.9593
10.1745 | -104.973
10.1782 | -104.9806
10.182 | -104.9857
10.1857 | -104.9911
10.1894 | -104.9968
10.1931 | -104.9996
10.1968 | -105.0014
10.2005 | -105.0042
10.2042 | -105.0053
10.2079 | -105.0052
10.2116 | -105.0029
10.2153 | -105.0024
10.219 | -104.9975
10.2227 | -104.9909
10.2264 | -104.9807
10.2301 | -104.9724
10.2338 | -104.9661
10.2375 | -104.9542
10.2412 | -104.943
10.2449 | -104.9322
10.2486 | -104.9246
10.2523 | -104.9178
10.256 | -104.9119
10.2597 | -104.9047
10.2634 | -104.897
10.2671 | -104.8901
10.2708 | -104.88
10.2745 | -104.8677
10.2783 | -104.858
10.282 | -104.8502
10.2857 | -104.8452
10.2894 | -104.8416
10.2931 | -104.8378
10.2968 | -104.8291
10.3005 | -104.8186
10.3042 | -104.8086
10.3079 | -104.8043
10.3116 | -104.7979
10.3153 | -104.7915
10.319 | -104.7851
10.3227 | -104.7799
10.3264 | -104.7774
10.3301 | -104.7769
10.3338 | -104.7781
10.3375 | -104.779
10.3412 | -104.7853
10.3449 | -104.797
10.3486 | -104.8097
10.3523 | -104.8245
10.356 | -104.8444
10.3597 | -104.8663
10.3634 | -104.894
10.3671 | -104.9195
10.3709 | -104.9456
10.3746 | -104.9705
10.3783 | -104.9958
10.382 | -105.0212
10.3857 | -105.0395
10.3894 | -105.0538
10.3931 | -105.0631
10.3968 | -105.0695
10.4005 | -105.0697
10.4042 | -105.065
10.4079 | -105.062
10.4116 | -105.0539
10.4153 | -105.0445
10.419 | -105.0329
10.4227 | -105.0202
10.4264 | -105.0078
10.4301 | -104.9938
10.4338 | -104.9767
10.4375 | -104.9619
10.4412 | -104.9484
10.4449 | -104.9382
10.4486 | -104.9295
10.4523 | -104.9211
10.456 | -104.9169
10.4597 | -104.9133
10.4634 | -104.9104
10.4672 | -104.9073
10.4709 | -104.9033
10.4746 | -104.902
10.4783 | -104.8975
10.482 | -104.8941
10.4857 | -104.8928
10.4894 | -104.8911
10.4931 | -104.8837
10.4968 | -104.8777
10.5005 | -104.8745
10.5042 | -104.8705
10.5079 | -104.8672
10.5116 | -104.8634
10.5153 | -104.86
10.519 | -104.8609
10.5227 | -104.8598
10.5264 | -104.8601
10.5301 | -104.8601
10.5338 | -104.8594
10.5375 | -104.8656
10.5412 | -104.8751
10.5449 | -104.8756
10.5486 | -104.8742
10.5523 | -104.874
10.556 | -104.8735
10.5597 | -104.8694
10.5635 | -104.8663
10.5672 | -104.8597
10.5709 | -104.8514
10.5746 | -104.841
10.5783 | -104.8299
10.582 | -104.8182
10.5857 | -104.8043
10.5894 | -104.7858
10.5931 | -104.7663
10.5968 | -104.7497
10.6005 | -104.7351
10.6042 | -104.7206
10.6079 | -104.7063
10.6116 | -104.6961
10.6153 | -104.689
10.619 | -104.6874
10.6227 | -104.6915
10.6264 | -104.6991
10.6301 | -104.7085
10.6338 | -104.7197
10.6375 | -104.7332
10.6412 | -104.7499
10.6449 | -104.77
10.6486 | -104.788
10.6523 | -104.8094
10.656 | -104.8297
10.6598 | -104.8457
10.6635 | -104.8617
10.6672 | -104.8774
10.6709 | -104.8923
10.6746 | -104.9134
10.6783 | -104.9334
10.682 | -104.952
10.6857 | -104.97
10.6894 | -104.9899
10.6931 | -105.0064
10.6968 | -105.0231
10.7005 | -105.0483
10.7042 | -105.0749
10.7079 | -105.1
10.7116 | -105.1234
10.7153 | -105.1443
10.719 | -105.1674
10.7227 | -105.1879
10.7264 | -105.2091
10.7301 | -105.23
10.7338 | -105.2482
10.7375 | -105.2665
10.7412 | -105.2814
10.7449 | -105.2954
10.7486 | -105.3128
10.7523 | -105.3274
10.7561 | -105.337
10.7598 | -105.3445
10.7635 | -105.3497
10.7672 | -105.3523
10.7709 | -105.3554
10.7746 | -105.351
10.7783 | -105.3449
10.782 | -105.3418
10.7857 | -105.3421
10.7894 | -105.3393
10.7931 | -105.3313
10.7968 | -105.3176
10.8005 | -105.3038
10.8042 | -105.29
10.8079 | -105.2695
10.8116 | -105.2503
10.8153 | -105.2333
10.819 | -105.2153
10.8227 | -105.1982
10.8264 | -105.1764
10.8301 | -105.1484
10.8338 | -105.117
10.8375 | -105.0873
10.8412 | -105.0586
10.8449 | -105.028
10.8486 | -104.9927
10.8524 | -104.9584
10.8561 | -104.9267
10.8598 | -104.8935
10.8635 | -104.8586
10.8672 | -104.8258
10.8709 | -104.7964
10.8746 | -104.7677
10.8783 | -104.7385
10.882 | -104.7159
10.8857 | -104.6931
10.8894 | -104.6697
10.8931 | -104.6518
10.8968 | -104.6383
10.9005 | -104.6323
10.9042 | -104.6268
10.9079 | -104.6238
10.9116 | -104.6249
10.9153 | -104.634
10.919 | -104.6491
10.9227 | -104.6691
10.9264 | -104.6924
10.9301 | -104.7183
10.9338 | -104.7507
10.9375 | -104.7839
10.9412 | -104.8164
10.945 | -104.8501
10.9487 | -104.8881
10.9524 | -104.9265
10.9561 | -104.9646
10.9598 | -104.9988
10.9635 | -105.0295
10.9672 | -105.0591
10.9709 | -105.0868
10.9746 | -105.112
10.9783 | -105.1309
10.982 | -105.1464
10.9857 | -105.1618
10.9894 | -105.1753
10.9931 | -105.1879
10.9968 | -105.1921
11.0005 | -105.1946
11.0042 | -105.1921
11.0079 | -105.1905
11.0116 | -105.1887
11.0153 | -105.184
11.019 | -105.1763
11.0227 | -105.1683
11.0264 | -105.1588
11.0301 | -105.1537
11.0338 | -105.1498
11.0375 | -105.1449
11.0413 | -105.141
11.045 | -105.1392
11.0487 | -105.1374
11.0524 | -105.1326
11.0561 | -105.1304
11.0598 | -105.1293
11.0635 | -105.1334
11.0672 | -105.1382
11.0709 | -105.1405
11.0746 | -105.1372
11.0783 | -105.136
11.082 | -105.1353
11.0857 | -105.1333
11.0894 | -105.1324
11.0931 | -105.1313
11.0968 | -105.1291
11.1005 | -105.1253
11.1042 | -105.1219
11.1079 | -105.1162
11.1116 | -105.1112
11.1153 | -105.1098
11.119 | -105.1065
11.1227 | -105.1026
11.1264 | -105.0944
11.1301 | -105.0824
11.1338 | -105.0673
11.1376 | -105.0532
11.1413 | -105.0382
11.145 | -105.0189
11.1487 | -105.0003
11.1524 | -104.9784
11.1561 | -104.9585
11.1598 | -104.94
11.1635 | -104.9223
11.1672 | -104.905
11.1709 | -104.8884
11.1746 | -104.8744
11.1783 | -104.8669
11.182 | -104.8611
11.1857 | -104.8594
11.1894 | -104.8565
11.1931 | -104.8566
11.1968 | -104.8608
11.2005 | -104.8702
11.2042 | -104.8821
11.2079 | -104.8991
11.2116 | -104.9177
11.2153 | -104.9342
11.219 | -104.9563
11.2227 | -104.9757
11.2264 | -104.9944
11.2302 | -105.0111
11.2339 | -105.031
11.2376 | -105.0544
11.2413 | -105.0779
11.245 | -105.0985
11.2487 | -105.1166
11.2524 | -105.1324
11.2561 | -105.1472
11.2598 | -105.1627
11.2635 | -105.1769
11.2672 | -105.1888
11.2709 | -105.1975
11.2746 | -105.2027
11.2783 | -105.2064
11.282 | -105.2051
11.2857 | -105.2028
11.2894 | -105.2029
11.2931 | -105.2011
11.2968 | -105.2002
11.3005 | -105.2002
11.3042 | -105.1984
11.3079 | -105.1994
11.3116 | -105.1966
11.3153 | -105.1962
11.319 | -105.1981
11.3227 | -105.1959
11.3264 | -105.1988
11.3302 | -105.2028
11.3339 | -105.2062
11.3376 | -105.2116
11.3413 | -105.2192
11.345 | -105.2331
11.3487 | -105.2469
11.3524 | -105.2642
11.3561 | -105.2812
11.3598 | -105.2995
11.3635 | -105.3181
11.3672 | -105.3392
11.3709 | -105.3635
11.3746 | -105.3911
11.3783 | -105.4201
11.382 | -105.4496
11.3857 | -105.4784
11.3894 | -105.5046
11.3931 | -105.5312
11.3968 | -105.5588
11.4005 | -105.5833
11.4042 | -105.6035
11.4079 | -105.6148
11.4116 | -105.6284
11.4153 | -105.6417
11.419 | -105.6477
11.4228 | -105.648
11.4265 | -105.6467
11.4302 | -105.6428
11.4339 | -105.6336
11.4376 | -105.6222
11.4413 | -105.6112
11.445 | -105.5937
11.4487 | -105.5787
11.4524 | -105.5558
11.4561 | -105.5254
11.4598 | -105.5006
11.4635 | -105.4739
11.4672 | -105.4429
11.4709 | -105.4128
11.4746 | -105.3865
11.4783 | -105.3624
11.482 | -105.3342
11.4857 | -105.3084
11.4894 | -105.2889
11.4931 | -105.2711
11.4968 | -105.2562
11.5005 | -105.2487
11.5042 | -105.2428
11.5079 | -105.2378
11.5116 | -105.2356
11.5153 | -105.2415
11.5191 | -105.252
11.5228 | -105.2664
11.5265 | -105.2838
11.5302 | -105.3016
11.5339 | -105.3201
11.5376 | -105.3358
11.5413 | -105.3517
11.545 | -105.3721
11.5487 | -105.3916
11.5524 | -105.4077
11.5561 | -105.425
11.5598 | -105.4418
11.5635 | -105.4551
11.5672 | -105.4665
11.5709 | -105.4787
11.5746 | -105.4917
11.5783 | -105.5014
11.582 | -105.511
11.5857 | -105.5217
11.5894 | -105.5262
11.5931 | -105.5278
11.5968 | -105.5286
11.6005 | -105.5278
11.6042 | -105.5253
11.6079 | -105.5206
11.6116 | -105.5185
11.6154 | -105.512
11.6191 | -105.5077
11.6228 | -105.5055
11.6265 | -105.4985
11.6302 | -105.4937
11.6339 | -105.4891
11.6376 | -105.4855
11.6413 | -105.4828
11.645 | -105.482
11.6487 | -105.4873
11.6524 | -105.4917
11.6561 | -105.4942
11.6598 | -105.4974
11.6635 | -105.5051
11.6672 | -105.5112
11.6709 | -105.518
11.6746 | -105.5234
11.6783 | -105.5279
11.682 | -105.5276
11.6857 | -105.5265
11.6894 | -105.5252
11.6931 | -105.5179
11.6968 | -105.51
11.7005 | -105.5053
11.7042 | -105.5
11.7079 | -105.4871
11.7117 | -105.4679
11.7154 | -105.4525
11.7191 | -105.4401
11.7228 | -105.4233
11.7265 | -105.4052
11.7302 | -105.3875
11.7339 | -105.3694
11.7376 | -105.3528
11.7413 | -105.337
11.745 | -105.3183
11.7487 | -105.2996
11.7524 | -105.2931
11.7561 | -105.2841
11.7598 | -105.2819
11.7635 | -105.2784
11.7672 | -105.2749
11.7709 | -105.2715
11.7746 | -105.2693
11.7783 | -105.2671
11.782 | -105.2688
11.7857 | -105.2761
11.7894 | -105.2842
11.7931 | -105.2928
11.7968 | -105.3056
11.8005 | -105.3173
11.8042 | -105.3295
11.808 | -105.3392
11.8117 | -105.3494
11.8154 | -105.3624
11.8191 | -105.3748
11.8228 | -105.3862
11.8265 | -105.3951
11.8302 | -105.402
11.8339 | -105.4052
11.8376 | -105.4104
11.8413 | -105.4169
11.845 | -105.4223
11.8487 | -105.4254
11.8524 | -105.4303
11.8561 | -105.4288
11.8598 | -105.4265
11.8635 | -105.4258
11.8672 | -105.4222
11.8709 | -105.4173
11.8746 | -105.4144
11.8783 | -105.4145
11.882 | -105.4112
11.8857 | -105.4043
11.8894 | -105.3945
11.8931 | -105.3858
11.8968 | -105.3776
11.9006 | -105.3688
11.9043 | -105.3607
11.908 | -105.3526
11.9117 | -105.3422
11.9154 | -105.3363
11.9191 | -105.3335
11.9228 | -105.3284
11.9265 | -105.3246
11.9302 | -105.3213
11.9339 | -105.318
11.9376 | -105.3177
11.9413 | -105.3216
11.945 | -105.3323
11.9487 | -105.3433
11.9524 | -105.3544
11.9561 | -105.3591
11.9598 | -105.3612
11.9635 | -105.3635
11.9672 | -105.3676
11.9709 | -105.3685
11.9746 | -105.3724
11.9783 | -105.3741
11.982 | -105.3745
11.9857 | -105.371
11.9894 | -105.3643
11.9931 | -105.3561
11.9969 | -105.3504
12.0006 | -105.3483
12.0043 | -105.3439
12.008 | -105.3351
12.0117 | -105.3279
12.0154 | -105.3218
12.0191 | -105.3171
12.0228 | -105.3132
12.0265 | -105.3088
12.0302 | -105.3074
12.0339 | -105.3077
12.0376 | -105.3096
12.0413 | -105.3111
12.045 | -105.3147
12.0487 | -105.3224
12.0524 | -105.3303
12.0561 | -105.338
12.0598 | -105.3515
12.0635 | -105.3667
12.0672 | -105.3808
12.0709 | -105.3974
12.0746 | -105.4146
12.0783 | -105.4317
12.082 | -105.4494
12.0857 | -105.4684
12.0894 | -105.4841
12.0932 | -105.4988
12.0969 | -105.5138
12.1006 | -105.5213
12.1043 | -105.5256
12.108 | -105.532
12.1117 | -105.5382
12.1154 | -105.5398
12.1191 | -105.5351
12.1228 | -105.5267
12.1265 | -105.5168
12.1302 | -105.5081
12.1339 | -105.501
12.1376 | -105.4936
12.1413 | -105.481
12.145 | -105.467
12.1487 | -105.4515
12.1524 | -105.4366
12.1561 | -105.4233
12.1598 | -105.4123
12.1635 | -105.397
12.1672 | -105.383
12.1709 | -105.3704
12.1746 | -105.3554
12.1783 | -105.3433
12.182 | -105.3333
12.1857 | -105.3194
12.1895 | -105.3029
12.1932 | -105.2896
12.1969 | -105.28
12.2006 | -105.2717
12.2043 | -105.2636
12.208 | -105.2548
12.2117 | -105.2452
12.2154 | -105.2399
12.2191 | -105.2335
12.2228 | -105.2256
12.2265 | -105.2195
12.2302 | -105.2181
12.2339 | -105.2124
12.2376 | -105.2074
12.2413 | -105.2079
12.245 | -105.2092
12.2487 | -105.2139
12.2524 | -105.2211
12.2561 | -105.228
12.2598 | -105.237
12.2635 | -105.2486
12.2672 | -105.2614
12.2709 | -105.2778
12.2746 | -105.2985
12.2783 | -105.3193
12.282 | -105.3418
12.2858 | -105.3633
12.2895 | -105.389
12.2932 | -105.4095
12.2969 | -105.4341
12.3006 | -105.4601
12.3043 | -105.4861
12.308 | -105.5108
12.3117 | -105.5322
12.3154 | -105.5534
12.3191 | -105.5754
12.3228 | -105.5968
12.3265 | -105.6131
12.3302 | -105.629
12.3339 | -105.6428
12.3376 | -105.6526
12.3413 | -105.6618
12.345 | -105.6678
12.3487 | -105.6708
12.3524 | -105.6725
12.3561 | -105.6744
12.3598 | -105.676
12.3635 | -105.6782
12.3672 | -105.6781
12.3709 | -105.6741
12.3746 | -105.6667
12.3784 | -105.6587
12.3821 | -105.6512
12.3858 | -105.6459
12.3895 | -105.6417
12.3932 | -105.6335
12.3969 | -105.624
12.4006 | -105.6165
12.4043 | -105.613
12.408 | -105.61
12.4117 | -105.6084
12.4154 | -105.6079
12.4191 | -105.6035
12.4228 | -105.6004
12.4265 | -105.6014
12.4302 | -105.6039
12.4339 | -105.6097
12.4376 | -105.6212
12.4413 | -105.637
12.445 | -105.6503
12.4487 | -105.6608
12.4524 | -105.6745
12.4561 | -105.6921
12.4598 | -105.7063
12.4635 | -105.7213
12.4672 | -105.739
12.4709 | -105.7543
12.4747 | -105.7694
12.4784 | -105.7825
12.4821 | -105.7915
12.4858 | -105.7967
12.4895 | -105.7994
12.4932 | -105.8018
12.4969 | -105.798
12.5006 | -105.7905
12.5043 | -105.7812
12.508 | -105.7731
12.5117 | -105.7622
12.5154 | -105.7465
12.5191 | -105.7314
12.5228 | -105.7189
12.5265 | -105.7027
12.5302 | -105.6831
12.5339 | -105.6646
12.5376 | -105.6434
12.5413 | -105.6219
12.545 | -105.6051
12.5487 | -105.5909
12.5524 | -105.5781
12.5561 | -105.5654
12.5598 | -105.5544
12.5635 | -105.5476
12.5672 | -105.5423
12.571 | -105.5404
12.5747 | -105.5416
12.5784 | -105.5467
12.5821 | -105.5524
12.5858 | -105.5587
12.5895 | -105.5638
12.5932 | -105.5708
12.5969 | -105.5794
12.6006 | -105.5941
12.6043 | -105.6095
12.608 | -105.6211
12.6117 | -105.6341
12.6154 | -105.6491
12.6191 | -105.6663
12.6228 | -105.6839
12.6265 | -105.7025
12.6302 | -105.7209
12.6339 | -105.7396
12.6376 | -105.7553
12.6413 | -105.7725
12.645 | -105.7908
12.6487 | -105.8032
12.6524 | -105.8173
12.6561 | -105.8281
12.6598 | -105.8344
12.6635 | -105.8427
12.6673 | -105.8498
12.671 | -105.8548
12.6747 | -105.8581
12.6784 | -105.8589
12.6821 | -105.8566
12.6858 | -105.8483
12.6895 | -105.8337
12.6932 | -105.8226
12.6969 | -105.8086
12.7006 | -105.7905
12.7043 | -105.7724
12.708 | -105.7541
12.7117 | -105.7372
12.7154 | -105.7229
12.7191 | -105.708
12.7228 | -105.6904
12.7265 | -105.667
12.7302 | -105.646
12.7339 | -105.6272
12.7376 | -105.6094
12.7413 | -105.5916
12.745 | -105.5717
12.7487 | -105.5524
12.7524 | -105.5325
12.7561 | -105.5155
12.7599 | -105.4999
12.7636 | -105.4839
12.7673 | -105.4649
12.771 | -105.4476
12.7747 | -105.4287
12.7784 | -105.4132
12.7821 | -105.3979
12.7858 | -105.3834
12.7895 | -105.3718
12.7932 | -105.3611
12.7969 | -105.3508
12.8006 | -105.3391
12.8043 | -105.3286
12.808 | -105.3216
12.8117 | -105.3186
12.8154 | -105.3141
12.8191 | -105.3091
12.8228 | -105.3109
12.8265 | -105.3047
12.8302 | -105.3011
12.8339 | -105.3009
12.8376 | -105.2999
12.8413 | -105.3026
12.845 | -105.3061
12.8487 | -105.3115
12.8524 | -105.3199
12.8562 | -105.3294
12.8599 | -105.3389
12.8636 | -105.3519
12.8673 | -105.3695
12.871 | -105.3861
12.8747 | -105.4024
12.8784 | -105.4155
12.8821 | -105.4312
12.8858 | -105.4507
12.8895 | -105.4656
12.8932 | -105.4797
12.8969 | -105.4894
12.9006 | -105.4937
12.9043 | -105.4985
12.908 | -105.5018
12.9117 | -105.5032
12.9154 | -105.5013
12.9191 | -105.4943
12.9228 | -105.4818
12.9265 | -105.4701
12.9302 | -105.4592
12.9339 | -105.4432
12.9376 | -105.4257
12.9413 | -105.4064
12.945 | -105.384
12.9487 | -105.3651
12.9525 | -105.3501
12.9562 | -105.3348
12.9599 | -105.3226
12.9636 | -105.3072
12.9673 | -105.2881
12.971 | -105.2748
12.9747 | -105.2654
12.9784 | -105.2561
12.9821 | -105.249
12.9858 | -105.2438
12.9895 | -105.2407
12.9932 | -105.2432
12.9969 | -105.2465
13.0006 | -105.2523
13.0043 | -105.2608
13.008 | -105.2682
13.0117 | -105.2773
13.0154 | -105.2886
13.0191 | -105.2985
13.0228 | -105.309
13.0265 | -105.3192
13.0302 | -105.3268
13.0339 | -105.3373
13.0376 | -105.3471
13.0413 | -105.3545
13.045 | -105.3587
13.0488 | -105.3643
13.0525 | -105.3734
13.0562 | -105.3797
13.0599 | -105.3879
13.0636 | -105.3944
13.0673 | -105.3964
13.071 | -105.3977
13.0747 | -105.3952
13.0784 | -105.3944
13.0821 | -105.3974
13.0858 | -105.4011
13.0895 | -105.4065
13.0932 | -105.4102
13.0969 | -105.4111
13.1006 | -105.4105
13.1043 | -105.4095
13.108 | -105.4093
13.1117 | -105.4108
13.1154 | -105.4092
13.1191 | -105.4059
13.1228 | -105.4049
13.1265 | -105.4041
13.1302 | -105.4041
13.1339 | -105.4064
13.1376 | -105.4055
13.1413 | -105.4018
13.1451 | -105.3973
13.1488 | -105.3916
13.1525 | -105.3901
13.1562 | -105.3922
13.1599 | -105.391
13.1636 | -105.3902
13.1673 | -105.3924
13.171 | -105.3926
13.1747 | -105.3958
13.1784 | -105.4003
13.1821 | -105.3999
13.1858 | -105.4015
13.1895 | -105.4047
13.1932 | -105.4103
13.1969 | -105.4189
13.2006 | -105.4299
13.2043 | -105.4408
13.208 | -105.4499
13.2117 | -105.4571
13.2154 | -105.4652
13.2191 | -105.4718
13.2228 | -105.4771
13.2265 | -105.4822
13.2302 | -105.4889
13.2339 | -105.4986
13.2377 | -105.5052
13.2414 | -105.5087
13.2451 | -105.5142
13.2488 | -105.5187
13.2525 | -105.5247
13.2562 | -105.5293
13.2599 | -105.5326
13.2636 | -105.5363
13.2673 | -105.5418
13.271 | -105.5434
13.2747 | -105.5434
13.2784 | -105.5423
13.2821 | -105.5406
13.2858 | -105.5404
13.2895 | -105.541
13.2932 | -105.5394
13.2969 | -105.5361
13.3006 | -105.5339
13.3043 | -105.5312
13.308 | -105.5288
13.3117 | -105.5244
13.3154 | -105.5178
13.3191 | -105.5077
13.3228 | -105.4982
13.3265 | -105.4908
13.3302 | -105.4856
13.334 | -105.481
13.3377 | -105.4738
13.3414 | -105.46
13.3451 | -105.4496
13.3488 | -105.4399
13.3525 | -105.4275
13.3562 | -105.4169
13.3599 | -105.4084
13.3636 | -105.4022
13.3673 | -105.3943
13.371 | -105.3872
13.3747 | -105.3813
13.3784 | -105.3704
13.3821 | -105.3607
13.3858 | -105.3545
13.3895 | -105.3454
13.3932 | -105.3361
13.3969 | -105.3292
13.4006 | -105.3252
13.4043 | -105.3239
13.408 | -105.3224
13.4117 | -105.3232
13.4154 | -105.3238
13.4191 | -105.3247
13.4228 | -105.3283
13.4265 | -105.3338
13.4303 | -105.3385
13.434 | -105.3426
13.4377 | -105.3479
13.4414 | -105.359
13.4451 | -105.3724
13.4488 | -105.3868
13.4525 | -105.397
13.4562 | -105.4087
13.4599 | -105.4213
13.4636 | -105.4358
13.4673 | -105.452
13.471 | -105.4704
13.4747 | -105.4888
13.4784 | -105.5044
13.4821 | -105.5205
13.4858 | -105.5379
13.4895 | -105.5528
13.4932 | -105.5633
13.4969 | -105.5737
13.5006 | -105.5836
13.5043 | -105.596
13.508 | -105.609
13.5117 | -105.6212
13.5154 | -105.6317
13.5191 | -105.6399
13.5228 | -105.6455
13.5266 | -105.6485
13.5303 | -105.6526
13.534 | -105.6613
13.5377 | -105.6642
13.5414 | -105.6628
13.5451 | -105.6634
13.5488 | -105.6613
13.5525 | -105.6588
13.5562 | -105.6581
13.5599 | -105.6551
13.5636 | -105.6488
13.5673 | -105.6417
13.571 | -105.6368
13.5747 | -105.6359
13.5784 | -105.6315
13.5821 | -105.6237
13.5858 | -105.6173
13.5895 | -105.6101
13.5932 | -105.6028
13.5969 | -105.596
13.6006 | -105.5887
13.6043 | -105.5822
13.608 | -105.5765
13.6117 | -105.5696
13.6154 | -105.5631
13.6191 | -105.5541
13.6229 | -105.5452
13.6266 | -105.5412
13.6303 | -105.5388
13.634 | -105.534
13.6377 | -105.5284
13.6414 | -105.5232
13.6451 | -105.5195
13.6488 | -105.5187
13.6525 | -105.5146
13.6562 | -105.5102
13.6599 | -105.5069
13.6636 | -105.5029
13.6673 | -105.4961
13.671 | -105.4906
13.6747 | -105.486
13.6784 | -105.4803
13.6821 | -105.4756
13.6858 | -105.4728
13.6895 | -105.4706
13.6932 | -105.4683
13.6969 | -105.4652
13.7006 | -105.4623
13.7043 | -105.4608
13.708 | -105.4603
13.7117 | -105.4591
13.7155 | -105.4595
13.7192 | -105.4639
13.7229 | -105.4716
13.7266 | -105.4772
13.7303 | -105.4831
13.734 | -105.4913
13.7377 | -105.5031
13.7414 | -105.5134
13.7451 | -105.5241
13.7488 | -105.5352
13.7525 | -105.5391
13.7562 | -105.5465
13.7599 | -105.5568
13.7636 | -105.5636
13.7673 | -105.5669
13.771 | -105.5705
13.7747 | -105.5733
13.7784 | -105.5744
13.7821 | -105.5746
13.7858 | -105.5713
13.7895 | -105.5685
13.7932 | -105.5657
13.7969 | -105.5628
13.8006 | -105.5589
13.8043 | -105.5518
13.808 | -105.5467
13.8118 | -105.5391
13.8155 | -105.5358
13.8192 | -105.5333
13.8229 | -105.5284
13.8266 | -105.5211
13.8303 | -105.5148
13.834 | -105.5086
13.8377 | -105.5056
13.8414 | -105.5029
13.8451 | -105.5031
13.8488 | -105.5074
13.8525 | -105.5105
13.8562 | -105.5145
13.8599 | -105.5245
13.8636 | -105.5375
13.8673 | -105.5502
13.871 | -105.5674
13.8747 | -105.5847
13.8784 | -105.601
13.8821 | -105.6195
13.8858 | -105.6409
13.8895 | -105.6645
13.8932 | -105.6853
13.8969 | -105.7056
13.9006 | -105.7262
13.9043 | -105.7467
13.9081 | -105.7673
13.9118 | -105.7877
13.9155 | -105.8077
13.9192 | -105.8255
13.9229 | -105.8437
13.9266 | -105.8621
13.9303 | -105.8766
13.934 | -105.8917
13.9377 | -105.9075
13.9414 | -105.9226
13.9451 | -105.9377
13.9488 | -105.9506
13.9525 | -105.9652
13.9562 | -105.9764
13.9599 | -105.9804
13.9636 | -105.9805
13.9673 | -105.9823
13.971 | -105.982
13.9747 | -105.981
13.9784 | -105.9757
13.9821 | -105.9657
13.9858 | -105.9592
13.9895 | -105.9473
13.9932 | -105.9294
13.9969 | -105.9083
14.0006 | -105.8882
14.0044 | -105.8711
14.0081 | -105.8469
14.0118 | -105.8197
14.0155 | -105.7893
14.0192 | -105.7576
14.0229 | -105.7271
14.0266 | -105.6971
14.0303 | -105.6663
14.034 | -105.6348
14.0377 | -105.5997
14.0414 | -105.5666
14.0451 | -105.5335
14.0488 | -105.5011
14.0525 | -105.473
14.0562 | -105.4477
14.0599 | -105.4242
14.0636 | -105.4005
14.0673 | -105.382
14.071 | -105.3668
14.0747 | -105.3568
14.0784 | -105.3497
14.0821 | -105.3434
14.0858 | -105.3399
14.0895 | -105.3439
14.0932 | -105.3482
14.0969 | -105.3559
14.1007 | -105.3685
14.1044 | -105.3841
14.1081 | -105.4001
14.1118 | -105.4191
14.1155 | -105.4395
14.1192 | -105.4592
14.1229 | -105.4803
14.1266 | -105.5014
14.1303 | -105.52
14.134 | -105.5403
14.1377 | -105.5603
14.1414 | -105.5786
14.1451 | -105.597
14.1488 | -105.6137
14.1525 | -105.6302
14.1562 | -105.6461
14.1599 | -105.6625
14.1636 | -105.681
14.1673 | -105.6995
14.171 | -105.7123
14.1747 | -105.7215
14.1784 | -105.7256
14.1821 | -105.7326
14.1858 | -105.7427
14.1895 | -105.7487
14.1933 | -105.7551
14.197 | -105.7619
14.2007 | -105.7704
14.2044 | -105.7778
14.2081 | -105.7814
14.2118 | -105.7906
14.2155 | -105.8042
14.2192 | -105.8161
14.2229 | -105.827
14.2266 | -105.8366
14.2303 | -105.8481
14.234 | -105.8616
14.2377 | -105.8753
14.2414 | -105.8899
14.2451 | -105.9062
14.2488 | -105.9224
14.2525 | -105.9355
14.2562 | -105.9477
14.2599 | -105.9583
14.2636 | -105.9674
14.2673 | -105.977
14.271 | -105.9835
14.2747 | -105.9911
14.2784 | -105.9926
14.2821 | -105.9931
14.2858 | -105.9957
14.2896 | -105.9969
14.2933 | -105.9914
14.297 | -105.9845
14.3007 | -105.9766
14.3044 | -105.9657
14.3081 | -105.9527
14.3118 | -105.9427
14.3155 | -105.9297
14.3192 | -105.9149
14.3229 | -105.8982
14.3266 | -105.8802
14.3303 | -105.8625
14.334 | -105.8474
14.3377 | -105.8308
14.3414 | -105.8127
14.3451 | -105.7949
14.3488 | -105.7787
14.3525 | -105.7651
14.3562 | -105.7532
14.3599 | -105.7438
14.3636 | -105.734
14.3673 | -105.7257
14.371 | -105.7213
14.3747 | -105.7207
14.3784 | -105.7189
14.3821 | -105.7225
14.3859 | -105.7294
14.3896 | -105.7367
14.3933 | -105.7448
14.397 | -105.7537
14.4007 | -105.7645
14.4044 | -105.7725
14.4081 | -105.7819
14.4118 | -105.7946
14.4155 | -105.8044
14.4192 | -105.8155
14.4229 | -105.8298
14.4266 | -105.8401
14.4303 | -105.8493
14.434 | -105.8602
14.4377 | -105.8713
14.4414 | -105.8804
14.4451 | -105.8895
14.4488 | -105.8988
14.4525 | -105.9069
14.4562 | -105.9146
14.4599 | -105.9245
14.4636 | -105.9319
14.4673 | -105.9385
14.471 | -105.9427
14.4747 | -105.9486
14.4784 | -105.9547
14.4822 | -105.9578
14.4859 | -105.9633
14.4896 | -105.9703
14.4933 | -105.9738
14.497 | -105.9754
14.5007 | -105.9786
14.5044 | -105.9806
14.5081 | -105.9791
14.5118 | -105.9796
14.5155 | -105.9805
14.5192 | -105.9793
14.5229 | -105.9783
14.5266 | -105.9802
14.5303 | -105.9845
14.534 | -105.9882
14.5377 | -105.9901
14.5414 | -105.9929
14.5451 | -105.9946
14.5488 | -105.9952
14.5525 | -105.9962
14.5562 | -105.9966
14.5599 | -105.9973
14.5636 | -105.998
14.5673 | -105.9985
14.571 | -105.9984
14.5747 | -105.9997
14.5785 | -105.9978
14.5822 | -105.9944
14.5859 | -105.9938
14.5896 | -105.9923
14.5933 | -105.9881
14.597 | -105.9845
14.6007 | -105.9792
14.6044 | -105.9734
14.6081 | -105.9677
14.6118 | -105.964
14.6155 | -105.9571
14.6192 | -105.947
14.6229 | -105.9365
14.6266 | -105.9213
14.6303 | -105.9053
14.634 | -105.8937
14.6377 | -105.8827
14.6414 | -105.8709
14.6451 | -105.8563
14.6488 | -105.8409
14.6525 | -105.8257
14.6562 | -105.8109
14.6599 | -105.7963
14.6636 | -105.7819
14.6673 | -105.7695
14.671 | -105.7583
14.6748 | -105.7444
14.6785 | -105.7314
14.6822 | -105.7235
14.6859 | -105.7164
14.6896 | -105.7085
14.6933 | -105.6974
14.697 | -105.6897
14.7007 | -105.6825
14.7044 | -105.6754
14.7081 | -105.6669
14.7118 | -105.6593
14.7155 | -105.6562
14.7192 | -105.6547
14.7229 | -105.6536
14.7266 | -105.6534
14.7303 | -105.6529
14.734 | -105.6508
14.7377 | -105.6502
14.7414 | -105.6496
14.7451 | -105.6474
14.7488 | -105.6465
14.7525 | -105.6463
14.7562 | -105.6465
14.7599 | -105.6473
14.7636 | -105.6481
14.7674 | -105.6485
14.7711 | -105.6486
14.7748 | -105.6471
14.7785 | -105.6465
14.7822 | -105.6458
14.7859 | -105.6419
14.7896 | -105.6399
14.7933 | -105.6369
14.797 | -105.6316
14.8007 | -105.6271
14.8044 | -105.6234
14.8081 | -105.6212
14.8118 | -105.6219
14.8155 | -105.6223
14.8192 | -105.6218
14.8229 | -105.6212
14.8266 | -105.6233
14.8303 | -105.6254
14.834 | -105.6294
14.8377 | -105.6361
14.8414 | -105.6427
14.8451 | -105.6485
14.8488 | -105.6565
14.8525 | -105.6666
14.8562 | -105.6796
14.8599 | -105.6943
14.8637 | -105.709
14.8674 | -105.725
14.8711 | -105.7415
14.8748 | -105.7581
14.8785 | -105.7752
14.8822 | -105.7911
14.8859 | -105.8117
14.8896 | -105.834
14.8933 | -105.8516
14.897 | -105.8653
14.9007 | -105.8794
14.9044 | -105.8981
14.9081 | -105.9122
14.9118 | -105.9239
14.9155 | -105.9335
14.9192 | -105.9423
14.9229 | -105.9487
14.9266 | -105.9502
14.9303 | -105.9482
14.934 | -105.9429
14.9377 | -105.9348
14.9414 | -105.9275
14.9451 | -105.9222
14.9488 | -105.9131
14.9525 | -105.9031
14.9562 | -105.894
14.96 | -105.8787
14.9637 | -105.8689
14.9674 | -105.8595
14.9711 | -105.8493
14.9748 | -105.8379
14.9785 | -105.8262
14.9822 | -105.8151
14.9859 | -105.8026
14.9896 | -105.7933
14.9933 | -105.785
14.997 | -105.7779
15.0007 | -105.7713
15.0044 | -105.762
15.0081 | -105.753
15.0118 | -105.747
15.0155 | -105.7419
15.0192 | -105.739
15.0229 | -105.7349
15.0266 | -105.7307
15.0303 | -105.7287
15.034 | -105.7272
15.0377 | -105.7254
15.0414 | -105.7215
15.0451 | -105.7192
15.0488 | -105.7157
15.0525 | -105.7091
15.0563 | -105.7007
15.06 | -105.6952
15.0637 | -105.6893
15.0674 | -105.6833
15.0711 | -105.6783
15.0748 | -105.6715
15.0785 | -105.6647
15.0822 | -105.66
15.0859 | -105.6553
15.0896 | -105.651
15.0933 | -105.6473
15.097 | -105.6453
15.1007 | -105.6407
15.1044 | -105.6353
15.1081 | -105.63
15.1118 | -105.6274
15.1155 | -105.6254
15.1192 | -105.6237
15.1229 | -105.6205
15.1266 | -105.6202
15.1303 | -105.6235
15.134 | -105.6255
15.1377 | -105.6266
15.1414 | -105.6289
15.1451 | -105.6325
15.1488 | -105.6356
15.1526 | -105.6368
15.1563 | -105.6401
15.16 | -105.6438
15.1637 | -105.6467
15.1674 | -105.6502
15.1711 | -105.655
15.1748 | -105.6569
15.1785 | -105.6592
15.1822 | -105.6614
15.1859 | -105.6637
15.1896 | -105.6658
15.1933 | -105.6673
15.197 | -105.669
15.2007 | -105.671
15.2044 | -105.6758
15.2081 | -105.6783
15.2118 | -105.6818
15.2155 | -105.6853
15.2192 | -105.6894
15.2229 | -105.6912
15.2266 | -105.6913
15.2303 | -105.6898
15.234 | -105.6854
15.2377 | -105.6777
15.2414 | -105.6742
15.2452 | -105.6737
15.2489 | -105.6711
15.2526 | -105.6699
15.2563 | -105.6684
15.26 | -105.6658
15.2637 | -105.6649
15.2674 | -105.6607
15.2711 | -105.6594
15.2748 | -105.6603
15.2785 | -105.654
15.2822 | -105.6478
15.2859 | -105.6441
15.2896 | -105.6395
15.2933 | -105.6343
15.297 | -105.6285
15.3007 | -105.6207
15.3044 | -105.6151
15.3081 | -105.6091
15.3118 | -105.6031
15.3155 | -105.5972
15.3192 | -105.5918
15.3229 | -105.5857
15.3266 | -105.5812
15.3303 | -105.5757
15.334 | -105.5703
15.3377 | -105.5636
15.3415 | -105.5587
15.3452 | -105.549
15.3489 | -105.5387
15.3526 | -105.5311
15.3563 | -105.5219
15.36 | -105.5132
15.3637 | -105.5054
15.3674 | -105.4989
15.3711 | -105.495
15.3748 | -105.488
15.3785 | -105.4804
15.3822 | -105.4749
15.3859 | -105.469
15.3896 | -105.4615
15.3933 | -105.4591
15.397 | -105.4541
15.4007 | -105.4504
15.4044 | -105.4473
15.4081 | -105.4478
15.4118 | -105.4484
15.4155 | -105.4462
15.4192 | -105.4452
15.4229 | -105.4432
15.4266 | -105.4426
15.4303 | -105.4427
15.434 | -105.4456
15.4378 | -105.4485
15.4415 | -105.4547
15.4452 | -105.4632
15.4489 | -105.4691
15.4526 | -105.4734
15.4563 | -105.4798
15.46 | -105.4836
15.4637 | -105.486
15.4674 | -105.491
15.4711 | -105.498
15.4748 | -105.5043
15.4785 | -105.5111
15.4822 | -105.5137
15.4859 | -105.5193
15.4896 | -105.5217
15.4933 | -105.5203
15.497 | -105.5187
15.5007 | -105.5204
15.5044 | -105.5225
15.5081 | -105.5195
15.5118 | -105.5126
15.5155 | -105.5092
15.5192 | -105.5064
15.5229 | -105.5008
15.5266 | -105.4936
15.5303 | -105.4868
15.5341 | -105.4784
15.5378 | -105.4701
15.5415 | -105.4632
15.5452 | -105.4532
15.5489 | -105.4428
15.5526 | -105.4322
15.5563 | -105.4191
15.56 | -105.405
15.5637 | -105.3919
15.5674 | -105.3832
15.5711 | -105.3746
15.5748 | -105.3649
15.5785 | -105.3561
15.5822 | -105.3479
15.5859 | -105.3401
15.5896 | -105.3317
15.5933 | -105.3262
15.597 | -105.3236
15.6007 | -105.3227
15.6044 | -105.324
15.6081 | -105.3274
15.6118 | -105.3335
15.6155 | -105.3417
15.6192 | -105.3514
15.6229 | -105.3599
15.6266 | -105.3694
15.6304 | -105.3782
15.6341 | -105.3897
15.6378 | -105.4047
15.6415 | -105.4212
15.6452 | -105.4387
15.6489 | -105.4572
15.6526 | -105.4732
15.6563 | -105.489
15.66 | -105.5089
15.6637 | -105.5291
15.6674 | -105.5449
15.6711 | -105.5621
15.6748 | -105.5804
15.6785 | -105.5965
15.6822 | -105.6121
15.6859 | -105.6262
15.6896 | -105.6373
15.6933 | -105.6461
15.697 | -105.6567
15.7007 | -105.6651
15.7044 | -105.6748
15.7081 | -105.6815
15.7118 | -105.6863
15.7155 | -105.6898
15.7192 | -105.6915
15.723 | -105.6925
15.7267 | -105.6918
15.7304 | -105.6873
15.7341 | -105.6853
15.7378 | -105.6848
15.7415 | -105.6801
15.7452 | -105.675
15.7489 | -105.6719
15.7526 | -105.6675
15.7563 | -105.6605
15.76 | -105.6546
15.7637 | -105.6517
15.7674 | -105.648
15.7711 | -105.6461
15.7748 | -105.6446
15.7785 | -105.6437
15.7822 | -105.6439
15.7859 | -105.6448
15.7896 | -105.6427
15.7933 | -105.6418
15.797 | -105.6433
15.8007 | -105.6472
15.8044 | -105.651
15.8081 | -105.6531
15.8118 | -105.6598
15.8155 | -105.6703
15.8193 | -105.6807
15.823 | -105.6892
15.8267 | -105.6979
15.8304 | -105.7068
15.8341 | -105.7176
15.8378 | -105.7284
15.8415 | -105.737
15.8452 | -105.7482
15.8489 | -105.7604
15.8526 | -105.7722
15.8563 | -105.7849
15.86 | -105.7951
15.8637 | -105.8022
15.8674 | -105.8045
15.8711 | -105.8071
15.8748 | -105.8079
15.8785 | -105.8107
15.8822 | -105.812
15.8859 | -105.8099
15.8896 | -105.8071
15.8933 | -105.8034
15.897 | -105.7972
15.9007 | -105.7884
15.9044 | -105.7769
15.9081 | -105.7658
15.9118 | -105.7522
15.9156 | -105.7382
15.9193 | -105.7252
15.923 | -105.7075
15.9267 | -105.6907
15.9304 | -105.6709
15.9341 | -105.6524
15.9378 | -105.6344
15.9415 | -105.6127
15.9452 | -105.5921
15.9489 | -105.5739
15.9526 | -105.5534
15.9563 | -105.5342
15.96 | -105.5163
15.9637 | -105.4967
15.9674 | -105.4778
15.9711 | -105.459
15.9748 | -105.439
15.9785 | -105.4207
15.9822 | -105.4044
15.9859 | -105.3901
15.9896 | -105.374
15.9933 | -105.3579
15.997 | -105.3413
16.0007 | -105.3289
16.0044 | -105.3172
16.0081 | -105.3071
16.0119 | -105.2996
16.0156 | -105.2949
16.0193 | -105.2876
16.023 | -105.278
16.0267 | -105.2684
16.0304 | -105.2625
16.0341 | -105.257
16.0378 | -105.2524
16.0415 | -105.247
16.0452 | -105.2407
16.0489 | -105.234
16.0526 | -105.23
16.0563 | -105.2283
16.06 | -105.2262
16.0637 | -105.2256
16.0674 | -105.2264
16.0711 | -105.2273
16.0748 | -105.2263
16.0785 | -105.226
16.0822 | -105.2253
16.0859 | -105.2283
16.0896 | -105.2321
16.0933 | -105.2358
16.097 | -105.2418
16.1007 | -105.2466
16.1044 | -105.2501
16.1082 | -105.2516
16.1119 | -105.2523
16.1156 | -105.2527
16.1193 | -105.2527
16.123 | -105.2538
16.1267 | -105.2541
16.1304 | -105.2525
16.1341 | -105.2505
16.1378 | -105.2462
16.1415 | -105.2404
16.1452 | -105.2337
16.1489 | -105.2243
16.1526 | -105.2148
16.1563 | -105.2041
16.16 | -105.1929
16.1637 | -105.1789
16.1674 | -105.1641
16.1711 | -105.1494
16.1748 | -105.1327
16.1785 | -105.1139
16.1822 | -105.0974
16.1859 | -105.0785
16.1896 | -105.0558
16.1933 | -105.0334
16.197 | -105.0114
16.2007 | -104.9911
16.2045 | -104.973
16.2082 | -104.9572
16.2119 | -104.9392
16.2156 | -104.9216
16.2193 | -104.9053
16.223 | -104.8899
16.2267 | -104.8777
16.2304 | -104.8627
16.2341 | -104.8481
16.2378 | -104.8353
16.2415 | -104.824
16.2452 | -104.8143
16.2489 | -104.807
16.2526 | -104.8017
16.2563 | -104.7986
16.26 | -104.7964
16.2637 | -104.7949
16.2674 | -104.7935
16.2711 | -104.7938
16.2748 | -104.7915
16.2785 | -104.7901
16.2822 | -104.7941
16.2859 | -104.8003
16.2896 | -104.8035
16.2933 | -104.8049
16.2971 | -104.8046
16.3008 | -104.8061
16.3045 | -104.8086
16.3082 | -104.8145
16.3119 | -104.8192
16.3156 | -104.8213
16.3193 | -104.822
16.323 | -104.8231
16.3267 | -104.8225
16.3304 | -104.8219
16.3341 | -104.8206
16.3378 | -104.8194
16.3415 | -104.8178
16.3452 | -104.8155
16.3489 | -104.8128
16.3526 | -104.8093
16.3563 | -104.8029
16.36 | -104.7969
16.3637 | -104.792
16.3674 | -104.7861
16.3711 | -104.7784
16.3748 | -104.7737
16.3785 | -104.7693
16.3822 | -104.7642
16.3859 | -104.761
16.3896 | -104.7583
16.3934 | -104.7576
16.3971 | -104.7593
16.4008 | -104.7617
16.4045 | -104.7655
16.4082 | -104.7714
16.4119 | -104.7783
16.4156 | -104.7859
16.4193 | -104.7984
16.423 | -104.8121
16.4267 | -104.8275
16.4304 | -104.843
16.4341 | -104.8585
16.4378 | -104.8767
16.4415 | -104.8946
16.4452 | -104.9162
16.4489 | -104.9353
16.4526 | -104.9578
16.4563 | -104.9853
16.46 | -105.0108
16.4637 | -105.0344
16.4674 | -105.0594
16.4711 | -105.0804
16.4748 | -105.0982
16.4785 | -105.1148
16.4822 | -105.1315
16.4859 | -105.1487
16.4897 | -105.163
16.4934 | -105.1763
16.4971 | -105.1876
16.5008 | -105.1951
16.5045 | -105.2007
16.5082 | -105.2047
16.5119 | -105.2057
16.5156 | -105.2024
16.5193 | -105.1949
16.523 | -105.1881
16.5267 | -105.1805
16.5304 | -105.17
16.5341 | -105.1523
16.5378 | -105.1358
16.5415 | -105.1152
16.5452 | -105.0945
16.5489 | -105.0757
16.5526 | -105.0538
16.5563 | -105.032
16.56 | -105.012
16.5637 | -104.9937
16.5674 | -104.9747
16.5711 | -104.9565
16.5748 | -104.9391
16.5785 | -104.9208
16.5822 | -104.902
16.586 | -104.8857
16.5897 | -104.8702
16.5934 | -104.8575
16.5971 | -104.8435
16.6008 | -104.8333
16.6045 | -104.8275
16.6082 | -104.8236
16.6119 | -104.8191
16.6156 | -104.8167
16.6193 | -104.8155
16.623 | -104.8157
16.6267 | -104.8175
16.6304 | -104.8218
16.6341 | -104.8264
16.6378 | -104.8307
16.6415 | -104.8379
16.6452 | -104.8452
16.6489 | -104.8533
16.6526 | -104.8622
16.6563 | -104.8707
16.66 | -104.8822
16.6637 | -104.8934
16.6674 | -104.9038
16.6711 | -104.9141
16.6748 | -104.9226
16.6786 | -104.9329
16.6823 | -104.9433
16.686 | -104.9549
16.6897 | -104.9642
16.6934 | -104.9715
16.6971 | -104.9819
16.7008 | -104.9947
16.7045 | -105.0065
16.7082 | -105.0199
16.7119 | -105.0333
16.7156 | -105.0477
16.7193 | -105.0594
16.723 | -105.0696
16.7267 | -105.0796
16.7304 | -105.09
16.7341 | -105.1039
16.7378 | -105.1166
16.7415 | -105.1284
16.7452 | -105.1431
16.7489 | -105.159
16.7526 | -105.1697
16.7563 | -105.181
16.76 | -105.1914
16.7637 | -105.2007
16.7674 | -105.2119
16.7711 | -105.2253
16.7749 | -105.2347
16.7786 | -105.2444
16.7823 | -105.2529
16.786 | -105.2614
16.7897 | -105.27
16.7934 | -105.2782
16.7971 | -105.2866
16.8008 | -105.2934
16.8045 | -105.2988
16.8082 | -105.3025
16.8119 | -105.3078
16.8156 | -105.314
16.8193 | -105.3188
16.823 | -105.3252
16.8267 | -105.3287
16.8304 | -105.3299
16.8341 | -105.33
16.8378 | -105.3338
16.8415 | -105.3344
16.8452 | -105.3332
16.8489 | -105.3318
16.8526 | -105.3327
16.8563 | -105.3323
16.86 | -105.3325
16.8637 | -105.3329
16.8674 | -105.3345
16.8712 | -105.336
16.8749 | -105.3378
16.8786 | -105.3422
16.8823 | -105.3451
16.886 | -105.3477
16.8897 | -105.3523
16.8934 | -105.3572
16.8971 | -105.3607
16.9008 | -105.3627
16.9045 | -105.3676
16.9082 | -105.3749
16.9119 | -105.3798
16.9156 | -105.3842
16.9193 | -105.3876
16.923 | -105.3925
16.9267 | -105.397
16.9304 | -105.4056
16.9341 | -105.4142
16.9378 | -105.4193
16.9415 | -105.4279
16.9452 | -105.4039
16.9489 | -105.3583
16.9526 | -105.3399
16.9563 | -105.3285
16.96 | -105.3157
16.9637 | -105.3025
16.9675 | -105.2875
16.9712 | -105.2668
16.9749 | -105.2416
16.9786 | -105.2148
16.9823 | -105.1857
16.986 | -105.1511
16.9897 | -105.1112
16.9934 | -105.0705
16.9971 | -105.0263
17.0008 | -104.9811
17.0045 | -104.934
17.0082 | -104.8829
17.0119 | -104.8302
17.0156 | -104.7763
17.0193 | -104.7206
17.023 | -104.6673
17.0267 | -104.6151
17.0304 | -104.5632
17.0341 | -104.5141
17.0378 | -104.4672
17.0415 | -104.4236
17.0452 | -104.3851
17.0489 | -104.3512
17.0526 | -104.3216
17.0563 | -104.2971
17.06 | -104.2786
17.0638 | -104.2648
17.0675 | -104.2564
17.0712 | -104.2552
17.0749 | -104.2581
17.0786 | -104.2642
17.0823 | -104.2739
17.086 | -104.2886
17.0897 | -104.3097
17.0934 | -104.3354
17.0971 | -104.3831
17.1008 | -104.4136
17.1045 | -104.4469
17.1082 | -104.4824
17.1119 | -104.5203
17.1156 | -104.5584
17.1193 | -104.5971
17.123 | -104.6371
17.1267 | -104.6793
17.1304 | -104.7223
17.1341 | -104.7634
17.1378 | -104.8045
17.1415 | -104.8466
17.1452 | -104.8883
17.1489 | -104.9266
17.1526 | -104.9634
17.1563 | -105.0001
17.1601 | -105.0351
17.1638 | -105.0678
17.1675 | -105.0987
17.1712 | -105.1299
17.1749 | -105.1548
17.1786 | -105.1806
17.1823 | -105.209
17.186 | -105.2328
17.1897 | -105.2818
17.1934 | -105.3386
17.1971 | -105.3605
17.2008 | -105.3708
17.2045 | -105.3806
17.2082 | -105.3891
17.2119 | -105.3986
17.2156 | -105.4105
17.2193 | -105.4219
17.223 | -105.435
17.2267 | -105.4461
17.2304 | -105.4566
17.2341 | -105.4707
17.2378 | -105.4847
17.2415 | -105.4948
17.2452 | -105.5037
17.2489 | -105.5128
17.2527 | -105.5237
17.2564 | -105.5332
17.2601 | -105.5434
17.2638 | -105.5542
17.2675 | -105.5606
17.2712 | -105.5654
17.2749 | -105.5691
17.2786 | -105.5719
17.2823 | -105.5748
17.286 | -105.5789
17.2897 | -105.5796
17.2934 | -105.5777
17.2971 | -105.5769
17.3008 | -105.5746
17.3045 | -105.573
17.3082 | -105.5704
17.3119 | -105.568
17.3156 | -105.5647
17.3193 | -105.561
17.323 | -105.5569
17.3267 | -105.5532
17.3304 | -105.5491
17.3341 | -105.5423
17.3378 | -105.5341
17.3415 | -105.5255
17.3452 | -105.5173
17.349 | -105.5103
17.3527 | -105.5046
17.3564 | -105.4964
17.3601 | -105.4877
17.3638 | -105.4785
17.3675 | -105.4719
17.3712 | -105.4647
17.3749 | -105.4551
17.3786 | -105.4455
17.3823 | -105.436
17.386 | -105.4269
17.3897 | -105.4182
17.3934 | -105.4083
17.3971 | -105.4005
17.4008 | -105.3915
17.4045 | -105.3853
17.4082 | -105.3779
17.4119 | -105.3713
17.4156 | -105.3656
17.4193 | -105.3556
17.423 | -105.3476
17.4267 | -105.3407
17.4304 | -105.3322
17.4341 | -105.3264
17.4378 | -105.3221
17.4415 | -105.3181
17.4453 | -105.3138
17.449 | -105.3124
17.4527 | -105.3125
17.4564 | -105.3126
17.4601 | -105.3121
17.4638 | -105.3122
17.4675 | -105.3145
17.4712 | -105.3145
17.4749 | -105.3166
17.4786 | -105.319
17.4823 | -105.318
17.486 | -105.3168
17.4897 | -105.3186
17.4934 | -105.3217
17.4971 | -105.324
17.5008 | -105.3254
17.5045 | -105.3277
17.5082 | -105.3314
17.5119 | -105.3346
17.5156 | -105.3385
17.5193 | -105.343
17.523 | -105.3505
17.5267 | -105.3546
17.5304 | -105.357
17.5341 | -105.3592
17.5378 | -105.3621
17.5416 | -105.3649
17.5453 | -105.3655
17.549 | -105.3654
17.5527 | -105.3636
17.5564 | -105.3602
17.5601 | -105.356
17.5638 | -105.3501
17.5675 | -105.3456
17.5712 | -105.3402
17.5749 | -105.3344
17.5786 | -105.329
17.5823 | -105.3246
17.586 | -105.3188
17.5897 | -105.3145
17.5934 | -105.3088
17.5971 | -105.3013
17.6008 | -105.2936
17.6045 | -105.2884
17.6082 | -105.2838
17.6119 | -105.2802
17.6156 | -105.2758
17.6193 | -105.2733
17.623 | -105.2712
17.6267 | -105.2693
17.6304 | -105.2752
17.6342 | -105.2784
17.6379 | -105.2818
17.6416 | -105.2854
17.6453 | -105.2919
17.649 | -105.2985
17.6527 | -105.3041
17.6564 | -105.3089
17.6601 | -105.315
17.6638 | -105.3215
17.6675 | -105.3295
17.6712 | -105.3414
17.6749 | -105.3525
17.6786 | -105.3617
17.6823 | -105.3721
17.686 | -105.3819
17.6897 | -105.3916
17.6934 | -105.4021
17.6971 | -105.412
17.7008 | -105.42
17.7045 | -105.428
17.7082 | -105.4353
17.7119 | -105.4425
17.7156 | -105.4496
17.7193 | -105.4558
17.723 | -105.4612
17.7267 | -105.4656
17.7305 | -105.4701
17.7342 | -105.4755
17.7379 | -105.4783
17.7416 | -105.4793
17.7453 | -105.4775
17.749 | -105.4759
17.7527 | -105.4692
17.7564 | -105.4601
17.7601 | -105.4482
17.7638 | -105.437
17.7675 | -105.4269
17.7712 | -105.4148
17.7749 | -105.4001
17.7786 | -105.388
17.7823 | -105.3761
17.786 | -105.3615
17.7897 | -105.346
17.7934 | -105.3284
17.7971 | -105.3092
17.8008 | -105.2924
17.8045 | -105.2774
17.8082 | -105.2624
17.8119 | -105.2455
17.8156 | -105.2301
17.8193 | -105.2161
17.823 | -105.1999
17.8268 | -105.1829
17.8305 | -105.1645
17.8342 | -105.1481
17.8379 | -105.1351
17.8416 | -105.1234
17.8453 | -105.112
17.849 | -105.0988
17.8527 | -105.0863
17.8564 | -105.077
17.8601 | -105.0686
17.8638 | -105.0626
17.8675 | -105.0551
17.8712 | -105.0468
17.8749 | -105.0413
17.8786 | -105.0378
17.8823 | -105.0361
17.886 | -105.034
17.8897 | -105.0295
17.8934 | -105.0253
17.8971 | -105.0234
17.9008 | -105.0216
17.9045 | -105.0219
17.9082 | -105.0246
17.9119 | -105.0272
17.9156 | -105.0313
17.9193 | -105.0339
17.9231 | -105.0345
17.9268 | -105.0351
17.9305 | -105.0373
17.9342 | -105.0402
17.9379 | -105.0419
17.9416 | -105.0433
17.9453 | -105.0464
17.949 | -105.0507
17.9527 | -105.0536
17.9564 | -105.0554
17.9601 | -105.0562
17.9638 | -105.0589
17.9675 | -105.0609
17.9712 | -105.0602
17.9749 | -105.0593
17.9786 | -105.0622
17.9823 | -105.0642
17.986 | -105.0656
17.9897 | -105.0666
17.9934 | -105.0676
17.9971 | -105.0669
18.0008 | -105.0656
18.0045 | -105.0639
18.0082 | -105.066
18.0119 | -105.0703
18.0156 | -105.0733
18.0194 | -105.0753
18.0231 | -105.0756
18.0268 | -105.0744
18.0305 | -105.0724
18.0342 | -105.07
18.0379 | -105.0688
18.0416 | -105.0659
18.0453 | -105.0631
18.049 | -105.0623
18.0527 | -105.0616
18.0564 | -105.057
18.0601 | -105.0496
18.0638 | -105.0435
18.0675 | -105.0389
18.0712 | -105.0335
18.0749 | -105.0249
18.0786 | -105.0159
18.0823 | -105.007
18.086 | -104.9982
18.0897 | -104.9877
18.0934 | -104.9775
18.0971 | -104.9664
18.1008 | -104.9543
18.1045 | -104.941
18.1082 | -104.9262
18.112 | -104.9104
18.1157 | -104.8924
18.1194 | -104.875
18.1231 | -104.8578
18.1268 | -104.8419
18.1305 | -104.8261
18.1342 | -104.8096
18.1379 | -104.792
18.1416 | -104.7756
18.1453 | -104.7614
18.149 | -104.746
18.1527 | -104.7302
18.1564 | -104.7144
18.1601 | -104.6969
18.1638 | -104.6815
18.1675 | -104.6737
18.1712 | -104.6591
18.1749 | -104.6448
18.1786 | -104.6336
18.1823 | -104.6219
18.186 | -104.6112
18.1897 | -104.5987
18.1934 | -104.5846
18.1971 | -104.5723
18.2008 | -104.5627
18.2045 | -104.5519
18.2082 | -104.5406
18.212 | -104.5316
18.2157 | -104.5238
18.2194 | -104.5157
18.2231 | -104.5061
18.2268 | -104.4976
18.2305 | -104.4902
18.2342 | -104.4817
18.2379 | -104.4719
18.2416 | -104.4627
18.2453 | -104.4545
18.249 | -104.4446
18.2527 | -104.4359
18.2564 | -104.4271
18.2601 | -104.4159
18.2638 | -104.4032
18.2675 | -104.3892
18.2712 | -104.3719
18.2749 | -104.3573
18.2786 | -104.3413
18.2823 | -104.3238
18.286 | -104.3089
18.2897 | -104.2953
18.2934 | -104.2772
18.2971 | -104.2548
18.3008 | -104.2344
18.3046 | -104.2146
18.3083 | -104.193
18.312 | -104.1725
18.3157 | -104.1546
18.3194 | -104.136
18.3231 | -104.1169
18.3268 | -104.0991
18.3305 | -104.0818
18.3342 | -104.0646
18.3379 | -104.0502
18.3416 | -104.0362
18.3453 | -104.0212
18.349 | -104.0068
18.3527 | -103.9932
18.3564 | -103.9801
18.3601 | -103.9695
18.3638 | -103.9607
18.3675 | -103.9527
18.3712 | -103.9435
18.3749 | -103.9342
18.3786 | -103.927
18.3823 | -103.9229
18.386 | -103.917
18.3897 | -103.9071
18.3934 | -103.8979
18.3971 | -103.8908
18.4009 | -103.8847
18.4046 | -103.8807
18.4083 | -103.8744
18.412 | -103.8691
18.4157 | -103.8633
18.4194 | -103.8572
18.4231 | -103.8501
18.4268 | -103.8429
18.4305 | -103.8344
18.4342 | -103.8268
18.4379 | -103.8183
18.4416 | -103.809
18.4453 | -103.7975
18.449 | -103.7832
18.4527 | -103.771
18.4564 | -103.7601
18.4601 | -103.7478
18.4638 | -103.7329
18.4675 | -103.7166
18.4712 | -103.6995
18.4749 | -103.6799
18.4786 | -103.659
18.4823 | -103.638
18.486 | -103.6173
18.4897 | -103.5961
18.4934 | -103.5715
18.4972 | -103.5467
18.5009 | -103.5231
18.5046 | -103.5
18.5083 | -103.477
18.512 | -103.4567
18.5157 | -103.4349
18.5194 | -103.4135
18.5231 | -103.3927
18.5268 | -103.3748
18.5305 | -103.3575
18.5342 | -103.3444
18.5379 | -103.3315
18.5416 | -103.3201
18.5453 | -103.3109
18.549 | -103.3035
18.5527 | -103.2971
18.5564 | -103.2935
18.5601 | -103.292
18.5638 | -103.2901
18.5675 | -103.2909
18.5712 | -103.2941
18.5749 | -103.2976
18.5786 | -103.3027
18.5823 | -103.3087
18.586 | -103.312
18.5897 | -103.3153
18.5935 | -103.3186
18.5972 | -103.3253
18.6009 | -103.3334
18.6046 | -103.3419
18.6083 | -103.3521
18.612 | -103.3625
18.6157 | -103.3708
18.6194 | -103.3798
18.6231 | -103.3897
18.6268 | -103.4
18.6305 | -103.4097
18.6342 | -103.4192
18.6379 | -103.4274
18.6416 | -103.4322
18.6453 | -103.4345
18.649 | -103.4389
18.6527 | -103.4441
18.6564 | -103.4474
18.6601 | -103.4489
18.6638 | -103.4492
18.6675 | -103.4471
18.6712 | -103.448
18.6749 | -103.4455
18.6786 | -103.4414
18.6823 | -103.4374
18.6861 | -103.433
18.6898 | -103.4278
18.6935 | -103.4253
18.6972 | -103.4216
18.7009 | -103.418
18.7046 | -103.4147
18.7083 | -103.4099
18.712 | -103.4061
18.7157 | -103.4063
18.7194 | -103.4041
18.7231 | -103.4011
18.7268 | -103.4
18.7305 | -103.3998
18.7342 | -103.401
18.7379 | -103.4039
18.7416 | -103.4094
18.7453 | -103.4141
18.749 | -103.4202
18.7527 | -103.4224
18.7564 | -103.4278
18.7601 | -103.4369
18.7638 | -103.4453
18.7675 | -103.4511
18.7712 | -103.4574
18.7749 | -103.4632
18.7786 | -103.468
18.7824 | -103.4735
18.7861 | -103.4807
18.7898 | -103.4887
18.7935 | -103.4948
18.7972 | -103.4996
18.8009 | -103.5014
18.8046 | -103.5027
18.8083 | -103.5055
18.812 | -103.5093
18.8157 | -103.512
18.8194 | -103.5137
18.8231 | -103.5166
18.8268 | -103.5194
18.8305 | -103.5211
18.8342 | -103.5229
18.8379 | -103.5258
18.8416 | -103.5274
18.8453 | -103.5283
18.849 | -103.5311
18.8527 | -103.535
18.8564 | -103.5369
18.8601 | -103.5397
18.8638 | -103.5394
18.8675 | -103.537
18.8712 | -103.5355
18.8749 | -103.5334
18.8787 | -103.5319
18.8824 | -103.5308
18.8861 | -103.5293
18.8898 | -103.5277
18.8935 | -103.5252
18.8972 | -103.5211
18.9009 | -103.5156
18.9046 | -103.5087
18.9083 | -103.5012
18.912 | -103.4935
18.9157 | -103.484
18.9194 | -103.4781
18.9231 | -103.4707
18.9268 | -103.4599
18.9305 | -103.4492
18.9342 | -103.4376
18.9379 | -103.4255
18.9416 | -103.4164
18.9453 | -103.4082
18.949 | -103.3979
18.9527 | -103.3877
18.9564 | -103.3767
18.9601 | -103.3657
18.9638 | -103.3545
18.9675 | -103.3447
18.9712 | -103.334
18.975 | -103.3215
18.9787 | -103.3085
18.9824 | -103.2969
18.9861 | -103.2837
18.9898 | -103.2725
18.9935 | -103.2607
18.9972 | -103.2474
19.0009 | -103.2345
19.0046 | -103.2232
19.0083 | -103.2117
19.012 | -103.2004
19.0157 | -103.19
19.0194 | -103.1768
19.0231 | -103.1649
19.0268 | -103.1501
19.0305 | -103.1361
19.0342 | -103.1238
19.0379 | -103.1121
19.0416 | -103.1003
19.0453 | -103.0887
19.049 | -103.0785
19.0527 | -103.07
19.0564 | -103.0607
19.0601 | -103.0497
19.0638 | -103.0388
19.0675 | -103.0277
19.0713 | -103.0179
19.075 | -103.0073
19.0787 | -102.996
19.0824 | -102.9838
19.0861 | -102.9706
19.0898 | -102.9597
19.0935 | -102.9489
19.0972 | -102.9388
19.1009 | -102.9284
19.1046 | -102.9137
19.1083 | -102.8963
19.112 | -102.8799
19.1157 | -102.8644
19.1194 | -102.848
19.1231 | -102.8344
19.1268 | -102.822
19.1305 | -102.8066
19.1342 | -102.791
19.1379 | -102.7746
19.1416 | -102.7566
19.1453 | -102.7381
19.149 | -102.7182
19.1527 | -102.7001
19.1564 | -102.6834
19.1601 | -102.6668
19.1639 | -102.6492
19.1676 | -102.633
19.1713 | -102.6218
19.175 | -102.614
19.1787 | -102.6067
19.1824 | -102.5994
19.1861 | -102.59
19.1898 | -102.5782
19.1935 | -102.5688
19.1972 | -102.5641
19.2009 | -102.5598
19.2046 | -102.5588
19.2083 | -102.5575
19.212 | -102.5548
19.2157 | -102.554
19.2194 | -102.554
19.2231 | -102.5536
19.2268 | -102.5561
19.2305 | -102.5597
19.2342 | -102.566
19.2379 | -102.5707
19.2416 | -102.575
19.2453 | -102.5815
19.249 | -102.5879
19.2527 | -102.5944
19.2564 | -102.6
19.2602 | -102.6039
19.2639 | -102.6064
19.2676 | -102.6114
19.2713 | -102.6165
19.275 | -102.621
19.2787 | -102.6253
19.2824 | -102.6284
19.2861 | -102.6315
19.2898 | -102.637
19.2935 | -102.6452
19.2972 | -102.6541
19.3009 | -102.6625
19.3046 | -102.6681
19.3083 | -102.6707
19.312 | -102.6775
19.3157 | -102.6855
19.3194 | -102.6937
19.3231 | -102.7008
19.3268 | -102.7094
19.3305 | -102.7194
19.3342 | -102.7312
19.3379 | -102.7415
19.3416 | -102.7534
19.3453 | -102.7651
19.349 | -102.7759
19.3527 | -102.7873
19.3565 | -102.8008
19.3602 | -102.8134
19.3639 | -102.8259
19.3676 | -102.8399
19.3713 | -102.8532
19.375 | -102.8675
19.3787 | -102.8838
19.3824 | -102.9009
19.3861 | -102.9184
19.3898 | -102.9334
19.3935 | -102.9475
19.3972 | -102.9603
19.4009 | -102.9729
19.4046 | -102.987
19.4083 | -103.0006
19.412 | -103.0129
19.4157 | -103.027
19.4194 | -103.0414
19.4231 | -103.0557
19.4268 | -103.066
19.4305 | -103.0742
19.4342 | -103.0823
19.4379 | -103.0921
19.4416 | -103.1011
19.4453 | -103.1065
19.449 | -103.111
19.4528 | -103.1149
19.4565 | -103.1188
19.4602 | -103.1236
19.4639 | -103.1294
19.4676 | -103.1353
19.4713 | -103.1379
19.475 | -103.1405
19.4787 | -103.1407
19.4824 | -103.1413
19.4861 | -103.1427
19.4898 | -103.144
19.4935 | -103.1458
19.4972 | -103.1473
19.5009 | -103.1468
19.5046 | -103.1456
19.5083 | -103.1472
19.512 | -103.1459
19.5157 | -103.1455
19.5194 | -103.1479
19.5231 | -103.1483
19.5268 | -103.1462
19.5305 | -103.1443
19.5342 | -103.1434
19.5379 | -103.1423
19.5416 | -103.1426
19.5453 | -103.143
19.5491 | -103.144
19.5528 | -103.1442
19.5565 | -103.1452
19.5602 | -103.1477
19.5639 | -103.1476
19.5676 | -103.1468
19.5713 | -103.1454
19.575 | -103.1465
19.5787 | -103.1482
19.5824 | -103.1502
19.5861 | -103.1541
19.5898 | -103.153
19.5935 | -103.1528
19.5972 | -103.1529
19.6009 | -103.1537
19.6046 | -103.1526
19.6083 | -103.1506
19.612 | -103.151
19.6157 | -103.1525
19.6194 | -103.1534
19.6231 | -103.1528
19.6268 | -103.152
19.6305 | -103.1511
19.6342 | -103.1456
19.6379 | -103.1417
19.6416 | -103.1385
19.6454 | -103.1355
19.6491 | -103.1335
19.6528 | -103.1321
19.6565 | -103.1289
19.6602 | -103.1262
19.6639 | -103.1221
19.6676 | -103.1193
19.6713 | -103.118
19.675 | -103.1148
19.6787 | -103.1103
19.6824 | -103.105
19.6861 | -103.1003
19.6898 | -103.0985
19.6935 | -103.0944
19.6972 | -103.0887
19.7009 | -103.0821
19.7046 | -103.0794
19.7083 | -103.0762
19.712 | -103.0714
19.7157 | -103.0664
19.7194 | -103.0631
19.7231 | -103.0601
19.7268 | -103.0593
19.7305 | -103.059
19.7342 | -103.0594
19.738 | -103.0595
19.7417 | -103.0605
19.7454 | -103.0622
19.7491 | -103.0661
19.7528 | -103.0705
19.7565 | -103.0765
19.7602 | -103.0839
19.7639 | -103.0926
19.7676 | -103.1129
19.7713 | -103.1221
19.775 | -103.1334
19.7787 | -103.1478
19.7824 | -103.1642
19.7861 | -103.1823
19.7898 | -103.2006
19.7935 | -103.2185
19.7972 | -103.2355
19.8009 | -103.2518
19.8046 | -103.2717
19.8083 | -103.2972
19.812 | -103.3236
19.8157 | -103.348
19.8194 | -103.3722
19.8231 | -103.3944
19.8268 | -103.4155
19.8305 | -103.4379
19.8343 | -103.461
19.838 | -103.4825
19.8417 | -103.5024
19.8454 | -103.5206
19.8491 | -103.538
19.8528 | -103.5567
19.8565 | -103.5708
19.8602 | -103.5822
19.8639 | -103.5941
19.8676 | -103.6015
19.8713 | -103.6089
19.875 | -103.6138
19.8787 | -103.6187
19.8824 | -103.6207
19.8861 | -103.6214
19.8898 | -103.6175
19.8935 | -103.6113
19.8972 | -103.6078
19.9009 | -103.6022
19.9046 | -103.5942
19.9083 | -103.5878
19.912 | -103.5792
19.9157 | -103.5706
19.9194 | -103.5607
19.9231 | -103.5481
19.9268 | -103.5322
19.9306 | -103.5149
19.9343 | -103.4971
19.938 | -103.4803
19.9417 | -103.4623
19.9454 | -103.4445
19.9491 | -103.4275
19.9528 | -103.4084
19.9565 | -103.3914
19.9602 | -103.3774
19.9639 | -103.3635
19.9676 | -103.351
19.9713 | -103.337
19.975 | -103.3237
19.9787 | -103.3113
19.9824 | -103.3002
19.9861 | -103.2874
19.9898 | -103.2772
19.9935 | -103.2698
19.9972 | -103.2635
20.0009 | -103.2566
20.0046 | -103.2499
20.0083 | -103.2433
20.012 | -103.2391
20.0157 | -103.2377
20.0194 | -103.238
20.0231 | -103.2375
20.0269 | -103.2373
20.0306 | -103.2389
20.0343 | -103.244
20.038 | -103.2501
20.0417 | -103.2575
20.0454 | -103.2659
20.0491 | -103.2744
20.0528 | -103.284
20.0565 | -103.2952
20.0602 | -103.3059
20.0639 | -103.3164
20.0676 | -103.327
20.0713 | -103.3394
20.075 | -103.3542
20.0787 | -103.3693
20.0824 | -103.3861
20.0861 | -103.4038
20.0898 | -103.4193
20.0935 | -103.4334
20.0972 | -103.4472
20.1009 | -103.4613
20.1046 | -103.4761
20.1083 | -103.4924
20.112 | -103.5062
20.1157 | -103.5195
20.1195 | -103.5324
20.1232 | -103.5455
20.1269 | -103.5601
20.1306 | -103.5752
20.1343 | -103.5887
20.138 | -103.599
20.1417 | -103.6108
20.1454 | -103.6239
20.1491 | -103.6346
20.1528 | -103.6465
20.1565 | -103.6566
20.1602 | -103.6643
20.1639 | -103.6702
20.1676 | -103.6761
20.1713 | -103.6792
20.175 | -103.6846
20.1787 | -103.6847
20.1824 | -103.6801
20.1861 | -103.6784
20.1898 | -103.6779
20.1935 | -103.6734
20.1972 | -103.6691
20.2009 | -103.6643
20.2046 | -103.6568
20.2083 | -103.6477
20.212 | -103.6386
20.2158 | -103.6316
20.2195 | -103.6242
20.2232 | -103.6149
20.2269 | -103.6058
20.2306 | -103.5967
20.2343 | -103.5866
20.238 | -103.5767
20.2417 | -103.5678
20.2454 | -103.5564
20.2491 | -103.5464
20.2528 | -103.5352
20.2565 | -103.5238
20.2602 | -103.5138
20.2639 | -103.5064
20.2676 | -103.499
20.2713 | -103.4908
20.275 | -103.4828
20.2787 | -103.4731
20.2824 | -103.4618
20.2861 | -103.4516
20.2898 | -103.4446
20.2935 | -103.4401
20.2972 | -103.4363
20.3009 | -103.4328
20.3046 | -103.4347
20.3083 | -103.4305
20.3121 | -103.429
20.3158 | -103.4275
20.3195 | -103.4255
20.3232 | -103.4234
20.3269 | -103.4225
20.3306 | -103.4215
20.3343 | -103.4214
20.338 | -103.422
20.3417 | -103.4239
20.3454 | -103.4279
20.3491 | -103.4319
20.3528 | -103.4344
20.3565 | -103.4368
20.3602 | -103.4388
20.3639 | -103.4433
20.3676 | -103.4497
20.3713 | -103.4548
20.375 | -103.4588
20.3787 | -103.4635
20.3824 | -103.4696
20.3861 | -103.4775
20.3898 | -103.4858
20.3935 | -103.494
20.3972 | -103.5008
20.4009 | -103.5091
20.4046 | -103.5196
20.4084 | -103.5298
20.4121 | -103.5398
20.4158 | -103.5485
20.4195 | -103.5594
20.4232 | -103.5703
20.4269 | -103.5837
20.4306 | -103.5971
20.4343 | -103.6078
20.438 | -103.6212
20.4417 | -103.6358
20.4454 | -103.6487
20.4491 | -103.6578
20.4528 | -103.6689
20.4565 | -103.6817
20.4602 | -103.6957
20.4639 | -103.7082
20.4676 | -103.7221
20.4713 | -103.7361
20.475 | -103.7466
20.4787 | -103.7524
20.4824 | -103.7599
20.4861 | -103.7695
20.4898 | -103.7782
20.4935 | -103.7873
20.4972 | -103.796
20.5009 | -103.8045
20.5047 | -103.8093
20.5084 | -103.8125
20.5121 | -103.8144
20.5158 | -103.8162
20.5195 | -103.8158
20.5232 | -103.8129
20.5269 | -103.8127
20.5306 | -103.8092
20.5343 | -103.8039
20.538 | -103.7994
20.5417 | -103.7954
20.5454 | -103.7897
20.5491 | -103.7819
20.5528 | -103.7731
20.5565 | -103.7661
20.5602 | -103.7582
20.5639 | -103.7457
20.5676 | -103.7329
20.5713 | -103.7201
20.575 | -103.7069
20.5787 | -103.6903
20.5824 | -103.6695
20.5861 | -103.6468
20.5898 | -103.6243
20.5935 | -103.6021
20.5972 | -103.5792
20.601 | -103.5558
20.6047 | -103.5311
20.6084 | -103.5076
20.6121 | -103.4825
20.6158 | -103.4553
20.6195 | -103.4272
20.6232 | -103.4002
20.6269 | -103.3718
20.6306 | -103.3424
20.6343 | -103.3123
20.638 | -103.2806
20.6417 | -103.2503
20.6454 | -103.2197
20.6491 | -103.1864
20.6528 | -103.1541
20.6565 | -103.1215
20.6602 | -103.0903
20.6639 | -103.059
20.6676 | -103.0267
20.6713 | -102.994
20.675 | -102.9617
20.6787 | -102.9286
20.6824 | -102.8971
20.6861 | -102.8678
20.6898 | -102.8374
20.6936 | -102.8084
20.6973 | -102.779
20.701 | -102.7493
20.7047 | -102.7217
20.7084 | -102.6951
20.7121 | -102.6697
20.7158 | -102.6449
20.7195 | -102.6177
20.7232 | -102.5914
20.7269 | -102.5667
20.7306 | -102.5462
20.7343 | -102.5229
20.738 | -102.4987
20.7417 | -102.4782
20.7454 | -102.4535
20.7491 | -102.4288
20.7528 | -102.4049
20.7565 | -102.3831
20.7602 | -102.3633
20.7639 | -102.3419
20.7676 | -102.3194
20.7713 | -102.3
20.775 | -102.2809
20.7787 | -102.256
20.7824 | -102.2315
20.7861 | -102.2093
20.7899 | -102.1852
20.7936 | -102.1607
20.7973 | -102.1376
20.801 | -102.1167
20.8047 | -102.0968
20.8084 | -102.075
20.8121 | -102.0511
20.8158 | -102.0279
20.8195 | -102.0037
20.8232 | -101.9801
20.8269 | -101.9573
20.8306 | -101.9318
20.8343 | -101.9092
20.838 | -101.8922
20.8417 | -101.8719
20.8454 | -101.8539
20.8491 | -101.8359
20.8528 | -101.8171
20.8565 | -101.8018
20.8602 | -101.79
20.8639 | -101.7791
20.8676 | -101.7671
20.8713 | -101.7556
20.875 | -101.7465
20.8787 | -101.7405
20.8824 | -101.7363
20.8862 | -101.7314
20.8899 | -101.7253
20.8936 | -101.7201
20.8973 | -101.7174
20.901 | -101.7155
20.9047 | -101.7145
20.9084 | -101.716
20.9121 | -101.7192
20.9158 | -101.724
20.9195 | -101.7288
20.9232 | -101.7357
20.9269 | -101.7423
20.9306 | -101.7499
20.9343 | -101.7583
20.938 | -101.7659
20.9417 | -101.7749
20.9454 | -101.7872
20.9491 | -101.7976
20.9528 | -101.8063
20.9565 | -101.8129
20.9602 | -101.8188
20.9639 | -101.8267
20.9676 | -101.8343
20.9713 | -101.8419
20.975 | -101.8467
20.9787 | -101.8466
20.9825 | -101.8441
20.9862 | -101.8452
20.9899 | -101.844
20.9936 | -101.8415
20.9973 | -101.8397
21.001 | -101.8359
21.0047 | -101.8318
21.0084 | -101.8275
21.0121 | -101.8215
21.0158 | -101.8173
21.0195 | -101.8128
21.0232 | -101.8083
21.0269 | -101.7983
21.0306 | -101.7899
21.0343 | -101.7838
21.038 | -101.7753
21.0417 | -101.7682
21.0454 | -101.7609
21.0491 | -101.7529
21.0528 | -101.7438
21.0565 | -101.7341
21.0602 | -101.7257
21.0639 | -101.719
21.0676 | -101.7131
21.0713 | -101.7064
21.0751 | -101.7027
21.0788 | -101.7008
21.0825 | -101.6976
21.0862 | -101.696
21.0899 | -101.695
21.0936 | -101.6944
21.0973 | -101.6928
21.101 | -101.6926
21.1047 | -101.6954
21.1084 | -101.7017
21.1121 | -101.708
21.1158 | -101.7133
21.1195 | -101.7182
21.1232 | -101.7248
21.1269 | -101.7343
21.1306 | -101.742
21.1343 | -101.7503
21.138 | -101.7623
21.1417 | -101.7747
21.1454 | -101.7849
21.1491 | -101.7967
21.1528 | -101.8094
21.1565 | -101.8219
21.1602 | -101.8332
21.1639 | -101.8429
21.1676 | -101.8545
21.1714 | -101.8643
21.1751 | -101.875
21.1788 | -101.8868
21.1825 | -101.8973
21.1862 | -101.9095
21.1899 | -101.9186
21.1936 | -101.9261
21.1973 | -101.9343
21.201 | -101.9418
21.2047 | -101.9485
21.2084 | -101.9528
21.2121 | -101.9571
21.2158 | -101.9593
21.2195 | -101.9615
21.2232 | -101.9607
21.2269 | -101.9602
21.2306 | -101.9557
21.2343 | -101.9497
21.238 | -101.9457
21.2417 | -101.937
21.2454 | -101.9233
21.2491 | -101.908
21.2528 | -101.8929
21.2565 | -101.876
21.2602 | -101.8578
21.2639 | -101.8394
21.2677 | -101.8197
21.2714 | -101.7974
21.2751 | -101.7761
21.2788 | -101.7543
21.2825 | -101.7316
21.2862 | -101.7067
21.2899 | -101.6812
21.2936 | -101.6528
21.2973 | -101.619
21.301 | -101.5852
21.3047 | -101.5513
21.3084 | -101.516
21.3121 | -101.4833
21.3158 | -101.4481
21.3195 | -101.4122
21.3232 | -101.3782
21.3269 | -101.3473
21.3306 | -101.3172
21.3343 | -101.2871
21.338 | -101.2566
21.3417 | -101.222
21.3454 | -101.1892
21.3491 | -101.1574
21.3528 | -101.1263
21.3565 | -101.0956
21.3602 | -101.0635
21.364 | -101.0326
21.3677 | -101.0022
21.3714 | -100.9809
21.3751 | -100.9519
21.3788 | -100.9222
21.3825 | -100.8938
21.3862 | -100.8679
21.3899 | -100.8442
21.3936 | -100.8202
21.3973 | -100.7961
21.401 | -100.7729
21.4047 | -100.7498
21.4084 | -100.7289
21.4121 | -100.7061
21.4158 | -100.6824
21.4195 | -100.6605
21.4232 | -100.6411
21.4269 | -100.6226
21.4306 | -100.6047
21.4343 | -100.5874
21.438 | -100.5704
21.4417 | -100.5516
21.4454 | -100.5343
21.4491 | -100.52
21.4528 | -100.5061
21.4565 | -100.4923
21.4603 | -100.4773
21.464 | -100.464
21.4677 | -100.4524
21.4714 | -100.4408
21.4751 | -100.4293
21.4788 | -100.4163
21.4825 | -100.4034
21.4862 | -100.3914
21.4899 | -100.3783
21.4936 | -100.3664
21.4973 | -100.3534
21.501 | -100.341
21.5047 | -100.3272
21.5084 | -100.3108
21.5121 | -100.2949
21.5158 | -100.2811
21.5195 | -100.2675
21.5232 | -100.255
21.5269 | -100.2425
21.5306 | -100.2299
21.5343 | -100.2159
21.538 | -100.2052
21.5417 | -100.1925
21.5454 | -100.1802
21.5491 | -100.1714
21.5528 | -100.1639
21.5566 | -100.1559
21.5603 | -100.1465
21.564 | -100.1346
21.5677 | -100.1211
21.5714 | -100.1078
21.5751 | -100.0934
21.5788 | -100.0804
21.5825 | -100.0681
21.5862 | -100.0538
21.5899 | -100.0388
21.5936 | -100.0264
21.5973 | -100.0156
21.601 | -100.0066
21.6047 | -99.9962
21.6084 | -99.9866
21.6121 | -99.9776
21.6158 | -99.9662
21.6195 | -99.9554
21.6232 | -99.9456
21.6269 | -99.9369
21.6306 | -99.9258
21.6343 | -99.9129
21.638 | -99.9025
21.6417 | -99.8946
21.6454 | -99.8901
21.6491 | -99.8852
21.6529 | -99.8778
21.6566 | -99.8685
21.6603 | -99.8593
21.664 | -99.8531
21.6677 | -99.8491
21.6714 | -99.8458
21.6751 | -99.8425
21.6788 | -99.8364
21.6825 | -99.8337
21.6862 | -99.8341
21.6899 | -99.835
21.6936 | -99.835
21.6973 | -99.8324
21.701 | -99.8318
21.7047 | -99.8329
21.7084 | -99.8337
21.7121 | -99.8357
21.7158 | -99.837
21.7195 | -99.84
21.7232 | -99.8451
21.7269 | -99.8511
21.7306 | -99.8549
21.7343 | -99.8578
21.738 | -99.8609
21.7417 | -99.8635
21.7455 | -99.8638
21.7492 | -99.8654
21.7529 | -99.8673
21.7566 | -99.8679
21.7603 | -99.8685
21.764 | -99.8704
21.7677 | -99.8725
21.7714 | -99.874
21.7751 | -99.8729
21.7788 | -99.872
21.7825 | -99.8717
21.7862 | -99.8673
21.7899 | -99.8616
21.7936 | -99.8597
21.7973 | -99.8606
21.801 | -99.8608
21.8047 | -99.8582
21.8084 | -99.8527
21.8121 | -99.8491
21.8158 | -99.8474
21.8195 | -99.8451
21.8232 | -99.8398
21.8269 | -99.8364
21.8306 | -99.8329
21.8343 | -99.8294
21.838 | -99.8266
21.8418 | -99.8248
21.8455 | -99.819
21.8492 | -99.8112
21.8529 | -99.8032
21.8566 | -99.7946
21.8603 | -99.7882
21.864 | -99.783
21.8677 | -99.7799
21.8714 | -99.7793
21.8751 | -99.777
21.8788 | -99.7747
21.8825 | -99.7736
21.8862 | -99.7691
21.8899 | -99.7661
21.8936 | -99.7654
21.8973 | -99.7635
21.901 | -99.7601
21.9047 | -99.7581
21.9084 | -99.7573
21.9121 | -99.7555
21.9158 | -99.7555
21.9195 | -99.7532
21.9232 | -99.7516
21.9269 | -99.7511
21.9306 | -99.7491
21.9343 | -99.748
21.9381 | -99.7471
21.9418 | -99.7492
21.9455 | -99.752
21.9492 | -99.7519
21.9529 | -99.7535
21.9566 | -99.7557
21.9603 | -99.7571
21.964 | -99.7604
21.9677 | -99.7671
21.9714 | -99.7734
21.9751 | -99.7797
21.9788 | -99.7856
21.9825 | -99.7914
21.9862 | -99.796
21.9899 | -99.8037
21.9936 | -99.8118
21.9973 | -99.8192
22.001 | -99.8255
22.0047 | -99.8333
22.0084 | -99.846
22.0121 | -99.86
22.0158 | -99.8731
22.0195 | -99.886
22.0232 | -99.8977
22.0269 | -99.9109
22.0307 | -99.9246
22.0344 | -99.9378
22.0381 | -99.9494
22.0418 | -99.964
22.0455 | -99.9788
22.0492 | -99.9927
22.0529 | -100.0062
22.0566 | -100.02
22.0603 | -100.0319
22.064 | -100.0419
22.0677 | -100.0472
22.0714 | -100.0509
22.0751 | -100.0493
22.0788 | -100.0509
22.0825 | -100.0523
22.0862 | -100.0557
22.0899 | -100.0585
22.0936 | -100.059
22.0973 | -100.0609
22.101 | -100.0601
22.1047 | -100.0547
22.1084 | -100.0449
22.1121 | -100.0367
22.1158 | -100.0299
22.1195 | -100.0193
22.1232 | -100.0056
22.127 | -99.9904
22.1307 | -99.9741
22.1344 | -99.9579
22.1381 | -99.9391
22.1418 | -99.9171
22.1455 | -99.8964
22.1492 | -99.8761
22.1529 | -99.8574
22.1566 | -99.8386
22.1603 | -99.8214
22.164 | -99.8057
22.1677 | -99.7882
22.1714 | -99.768
22.1751 | -99.7481
22.1788 | -99.7297
22.1825 | -99.7106
22.1862 | -99.6912
22.1899 | -99.672
22.1936 | -99.6558
22.1973 | -99.6434
22.201 | -99.6305
22.2047 | -99.6192
22.2084 | -99.6103
22.2121 | -99.6018
22.2158 | -99.5929
22.2195 | -99.5864
22.2233 | -99.584
22.227 | -99.5812
22.2307 | -99.5788
22.2344 | -99.5798
22.2381 | -99.5831
22.2418 | -99.5898
22.2455 | -99.6003
22.2492 | -99.6105
22.2529 | -99.6203
22.2566 | -99.6316
22.2603 | -99.6453
22.264 | -99.6603
22.2677 | -99.6747
22.2714 | -99.6916
22.2751 | -99.7107
22.2788 | -99.7316
22.2825 | -99.7502
22.2862 | -99.7698
22.2899 | -99.7912
22.2936 | -99.8112
22.2973 | -99.8313
22.301 | -99.8536
22.3047 | -99.8754
22.3084 | -99.8984
22.3121 | -99.9233
22.3158 | -99.9482
22.3196 | -99.9723
22.3233 | -99.9936
22.327 | -100.015
22.3307 | -100.0373
22.3344 | -100.0612
22.3381 | -100.0863
22.3418 | -100.1097
22.3455 | -100.1339
22.3492 | -100.1583
22.3529 | -100.182
22.3566 | -100.2048
22.3603 | -100.228
22.364 | -100.2514
22.3677 | -100.2763
22.3714 | -100.3029
22.3751 | -100.3285
22.3788 | -100.3532
22.3825 | -100.3812
22.3862 | -100.4076
22.3899 | -100.4375
22.3936 | -100.4614
22.3973 | -100.4825
22.401 | -100.5027
22.4047 | -100.522
22.4084 | -100.5412
22.4121 | -100.5608
22.4159 | -100.5816
22.4196 | -100.5998
22.4233 | -100.6194
22.427 | -100.638
22.4307 | -100.6573
22.4344 | -100.6787
22.4381 | -100.6992
22.4418 | -100.7139
22.4455 | -100.7334
22.4492 | -100.753
22.4529 | -100.7713
22.4566 | -100.7891
22.4603 | -100.8089
22.464 | -100.8265
22.4677 | -100.8408
22.4714 | -100.8542
22.4751 | -100.8682
22.4788 | -100.8819
22.4825 | -100.8949
22.4862 | -100.9073
22.4899 | -100.9205
22.4936 | -100.9336
22.4973 | -100.9452
22.501 | -100.9584
22.5047 | -100.9728
22.5084 | -100.9863
22.5122 | -100.9965
22.5159 | -101.0052
22.5196 | -101.0152
22.5233 | -101.0232
22.527 | -101.0278
22.5307 | -101.0337
22.5344 | -101.0401
22.5381 | -101.0445
22.5418 | -101.0472
22.5455 | -101.0503
22.5492 | -101.052
22.5529 | -101.052
22.5566 | -101.0482
22.5603 | -101.0428
22.564 | -101.0376
22.5677 | -101.0321
22.5714 | -101.0281
22.5751 | -101.0217
22.5788 | -101.0111
22.5825 | -101.004
22.5862 | -100.9979
22.5899 | -100.9908
22.5936 | -100.982
22.5973 | -100.9706
22.601 | -100.9604
22.6047 | -100.9507
22.6085 | -100.9365
22.6122 | -100.9243
22.6159 | -100.9137
22.6196 | -100.9021
22.6233 | -100.8897
22.627 | -100.8741
22.6307 | -100.8575
22.6344 | -100.8418
22.6381 | -100.8259
22.6418 | -100.8102
22.6455 | -100.7934
22.6492 | -100.7776
22.6529 | -100.7607
22.6566 | -100.7428
22.6603 | -100.7261
22.664 | -100.709
22.6677 | -100.6903
22.6714 | -100.6728
22.6751 | -100.6563
22.6788 | -100.6402
22.6825 | -100.6256
22.6862 | -100.61
22.6899 | -100.595
22.6936 | -100.5804
22.6973 | -100.567
22.7011 | -100.5553
22.7048 | -100.5443
22.7085 | -100.5293
22.7122 | -100.5141
22.7159 | -100.504
22.7196 | -100.4939
22.7233 | -100.482
22.727 | -100.4708
22.7307 | -100.4606
22.7344 | -100.4512
22.7381 | -100.446
22.7418 | -100.4402
22.7455 | -100.4332
22.7492 | -100.4266
22.7529 | -100.4197
22.7566 | -100.4141
22.7603 | -100.4095
22.764 | -100.4061
22.7677 | -100.4019
22.7714 | -100.3966
22.7751 | -100.3907
22.7788 | -100.3843
22.7825 | -100.3759
22.7862 | -100.3661
22.7899 | -100.3561
22.7936 | -100.3472
22.7974 | -100.3362
22.8011 | -100.326
22.8048 | -100.3153
22.8085 | -100.3038
22.8122 | -100.289
22.8159 | -100.274
22.8196 | -100.2623
22.8233 | -100.2505
22.827 | -100.2378
22.8307 | -100.2248
22.8344 | -100.2098
22.8381 | -100.1924
22.8418 | -100.1728
22.8455 | -100.1553
22.8492 | -100.1402
22.8529 | -100.1218
22.8566 | -100.1028
22.8603 | -100.0861
22.864 | -100.0703
22.8677 | -100.0521
22.8714 | -100.0336
22.8751 | -100.0169
22.8788 | -99.9998
22.8825 | -99.9814
22.8862 | -99.9627
22.8899 | -99.9445
22.8937 | -99.9267
22.8974 | -99.9099
22.9011 | -99.8938
22.9048 | -99.8743
22.9085 | -99.8555
22.9122 | -99.8389
22.9159 | -99.8225
22.9196 | -99.8077
22.9233 | -99.7945
22.927 | -99.7819
22.9307 | -99.7696
22.9344 | -99.7583
22.9381 | -99.7479
22.9418 | -99.7373
22.9455 | -99.7295
22.9492 | -99.7223
22.9529 | -99.7143
22.9566 | -99.707
22.9603 | -99.7014
22.964 | -99.6966
22.9677 | -99.6921
22.9714 | -99.688
22.9751 | -99.688
22.9788 | -99.686
22.9825 | -99.6841
22.9862 | -99.6786
22.99 | -99.674
22.9937 | -99.6728
22.9974 | -99.6726
23.0011 | -99.6713
23.0048 | -99.6676
23.0085 | -99.6641
23.0122 | -99.6617
23.0159 | -99.6594
23.0196 | -99.6594
23.0233 | -99.6584
23.027 | -99.6549
23.0307 | -99.6518
23.0344 | -99.6508
23.0381 | -99.6522
23.0418 | -99.652
23.0455 | -99.6508
23.0492 | -99.649
23.0529 | -99.6471
23.0566 | -99.6431
23.0603 | -99.6393
23.064 | -99.6371
23.0677 | -99.6348
23.0714 | -99.6341
23.0751 | -99.6328
23.0788 | -99.6323
23.0826 | -99.6316
23.0863 | -99.6315
23.09 | -99.6299
23.0937 | -99.6295
23.0974 | -99.6301
23.1011 | -99.6315
23.1048 | -99.6346
23.1085 | -99.6369
23.1122 | -99.6397
23.1159 | -99.6442
23.1196 | -99.6466
23.1233 | -99.648
23.127 | -99.6485
23.1307 | -99.6514
23.1344 | -99.6563
23.1381 | -99.6628
23.1418 | -99.6671
23.1455 | -99.6694
23.1492 | -99.6721
23.1529 | -99.6748
23.1566 | -99.6785
23.1603 | -99.6834
23.164 | -99.6901
23.1677 | -99.6963
23.1714 | -99.6994
23.1751 | -99.7035
23.1789 | -99.7109
23.1826 | -99.7186
23.1863 | -99.7249
23.19 | -99.7313
23.1937 | -99.7407
23.1974 | -99.7499
23.2011 | -99.7568
23.2048 | -99.7649
23.2085 | -99.7743
23.2122 | -99.7825
23.2159 | -99.7901
23.2196 | -99.7971
23.2233 | -99.8052
23.227 | -99.8157
23.2307 | -99.8268
23.2344 | -99.8349
23.2381 | -99.8406
23.2418 | -99.8484
23.2455 | -99.8562
23.2492 | -99.865
23.2529 | -99.8728
23.2566 | -99.8813
23.2603 | -99.8883
23.264 | -99.8945
23.2677 | -99.9016
23.2714 | -99.9071
23.2752 | -99.9129
23.2789 | -99.9197
23.2826 | -99.9265
23.2863 | -99.934
23.29 | -99.942
23.2937 | -99.949
23.2974 | -99.9558
23.3011 | -99.9627
23.3048 | -99.9693
23.3085 | -99.9742
23.3122 | -99.9819
23.3159 | -99.9937
23.3196 | -100.0029
23.3233 | -100.0075
23.327 | -100.0122
23.3307 | -100.0178
23.3344 | -100.0239
23.3381 | -100.029
23.3418 | -100.0331
23.3455 | -100.0376
23.3492 | -100.0423
23.3529 | -100.0463
23.3566 | -100.0535
23.3603 | -100.0608
23.364 | -100.0674
23.3677 | -100.0715
23.3715 | -100.0748
23.3752 | -100.0792
23.3789 | -100.0872
23.3826 | -100.0949
23.3863 | -100.104
23.39 | -100.1107
23.3937 | -100.1141
23.3974 | -100.1191
23.4011 | -100.1233
23.4048 | -100.1274
23.4085 | -100.1305
23.4122 | -100.1327
23.4159 | -100.1348
23.4196 | -100.1359
23.4233 | -100.137
23.427 | -100.1397
23.4307 | -100.1402
23.4344 | -100.14
23.4381 | -100.1399
23.4418 | -100.1383
23.4455 | -100.1351
23.4492 | -100.1345
23.4529 | -100.1357
23.4566 | -100.1357
23.4603 | -100.1335
23.464 | -100.1299
23.4678 | -100.1268
23.4715 | -100.1253
23.4752 | -100.1235
23.4789 | -100.1218
23.4826 | -100.1219
23.4863 | -100.1211
23.49 | -100.1207
23.4937 | -100.1206
23.4974 | -100.1212
23.5011 | -100.1234
23.5048 | -100.1236
23.5085 | -100.1235
23.5122 | -100.1221
23.5159 | -100.1213
23.5196 | -100.1213
23.5233 | -100.1202
23.527 | -100.1173
23.5307 | -100.1161
23.5344 | -100.1171
23.5381 | -100.1163
23.5418 | -100.1149
23.5455 | -100.1136
23.5492 | -100.1134
23.5529 | -100.1141
23.5566 | -100.1152
23.5603 | -100.1151
23.5641 | -100.1145
23.5678 | -100.1154
23.5715 | -100.1159
23.5752 | -100.1171
23.5789 | -100.1186
23.5826 | -100.1192
23.5863 | -100.1213
23.59 | -100.1246
23.5937 | -100.1271
23.5974 | -100.1305
23.6011 | -100.1342
23.6048 | -100.1373
23.6085 | -100.1407
23.6122 | -100.1447
23.6159 | -100.1488
23.6196 | -100.1515
23.6233 | -100.1539
23.627 | -100.1591
23.6307 | -100.1662
23.6344 | -100.1738
23.6381 | -100.182
23.6418 | -100.191
23.6455 | -100.1981
23.6492 | -100.2047
23.6529 | -100.2143
23.6567 | -100.2255
23.6604 | -100.2331
23.6641 | -100.2394
23.6678 | -100.2506
23.6715 | -100.2608
23.6752 | -100.2703
23.6789 | -100.2798
23.6826 | -100.2865
23.6863 | -100.2941
23.69 | -100.3
23.6937 | -100.3067
23.6974 | -100.3148
23.7011 | -100.3234
23.7048 | -100.3315
23.7085 | -100.3397
23.7122 | -100.3457
23.7159 | -100.3536
23.7196 | -100.3617
23.7233 | -100.3692
23.727 | -100.3752
23.7307 | -100.3812
23.7344 | -100.3877
23.7381 | -100.3938
23.7418 | -100.3992
23.7455 | -100.404
23.7492 | -100.4071
23.753 | -100.4112
23.7567 | -100.4141
23.7604 | -100.4154
23.7641 | -100.4177
23.7678 | -100.4197
23.7715 | -100.4227
23.7752 | -100.4256
23.7789 | -100.4262
23.7826 | -100.4243
23.7863 | -100.4207
23.79 | -100.418
23.7937 | -100.4129
23.7974 | -100.4087
23.8011 | -100.4073
23.8048 | -100.4058
23.8085 | -100.4035
23.8122 | -100.4003
23.8159 | -100.3972
23.8196 | -100.3937
23.8233 | -100.3891
23.827 | -100.383
23.8307 | -100.3789
23.8344 | -100.3758
23.8381 | -100.3732
23.8418 | -100.3709
23.8455 | -100.3687
23.8493 | -100.366
23.853 | -100.3631
23.8567 | -100.3591
23.8604 | -100.3566
23.8641 | -100.3555
23.8678 | -100.3517
23.8715 | -100.3484
23.8752 | -100.3462
23.8789 | -100.3422
23.8826 | -100.339
23.8863 | -100.3355
23.89 | -100.3303
23.8937 | -100.3255
23.8974 | -100.3221
23.9011 | -100.3171
23.9048 | -100.3146
23.9085 | -100.31
23.9122 | -100.3025
23.9159 | -100.2954
23.9196 | -100.2887
23.9233 | -100.2816
23.927 | -100.2761
23.9307 | -100.2688
23.9344 | -100.2577
23.9381 | -100.2482
23.9418 | -100.24
23.9456 | -100.2325
23.9493 | -100.2265
23.953 | -100.2209
23.9567 | -100.216
23.9604 | -100.2103
23.9641 | -100.2027
23.9678 | -100.194
23.9715 | -100.1855
23.9752 | -100.1777
23.9789 | -100.1701
23.9826 | -100.1623
23.9863 | -100.1551
23.99 | -100.1473
23.9937 | -100.1401
23.9974 | -100.135
24.0011 | -100.1289
24.0048 | -100.1188
24.0085 | -100.1098
24.0122 | -100.1005
24.0159 | -100.0901
24.0196 | -100.0826
24.0233 | -100.075
24.027 | -100.067
24.0307 | -100.0585
24.0344 | -100.0501
24.0382 | -100.0415
24.0419 | -100.0326
24.0456 | -100.0259
24.0493 | -100.0172
24.053 | -100.0081
24.0567 | -99.9987
24.0604 | -99.9896
24.0641 | -99.9817
24.0678 | -99.9744
24.0715 | -99.9656
24.0752 | -99.9581
24.0789 | -99.9514
24.0826 | -99.9454
24.0863 | -99.94
24.09 | -99.9343
24.0937 | -99.93
24.0974 | -99.9238
24.1011 | -99.9162
24.1048 | -99.9113
24.1085 | -99.9064
24.1122 | -99.9
24.1159 | -99.8944
24.1196 | -99.8868
24.1233 | -99.8792
24.127 | -99.8746
24.1307 | -99.8694
24.1345 | -99.8657
24.1382 | -99.8617
24.1419 | -99.8569
24.1456 | -99.8536
24.1493 | -99.8521
24.153 | -99.8512
24.1567 | -99.8493
24.1604 | -99.8476
24.1641 | -99.8477
24.1678 | -99.8499
24.1715 | -99.8497
24.1752 | -99.8479
24.1789 | -99.8454
24.1826 | -99.8435
24.1863 | -99.843
24.19 | -99.844
24.1937 | -99.8441
24.1974 | -99.8454
24.2011 | -99.8464
24.2048 | -99.8463
24.2085 | -99.848
24.2122 | -99.8501
24.2159 | -99.854
24.2196 | -99.8598
24.2233 | -99.8656
24.227 | -99.8702
24.2308 | -99.8729
24.2345 | -99.8742
24.2382 | -99.8759
24.2419 | -99.8788
24.2456 | -99.8795
24.2493 | -99.8827
24.253 | -99.8871
24.2567 | -99.8907
24.2604 | -99.8933
24.2641 | -99.8966
24.2678 | -99.8997
24.2715 | -99.9029
24.2752 | -99.9097
24.2789 | -99.9151
24.2826 | -99.9208
24.2863 | -99.9243
24.29 | -99.924
24.2937 | -99.9228
24.2974 | -99.9215
24.3011 | -99.9217
24.3048 | -99.9228
24.3085 | -99.9225
24.3122 | -99.9237
24.3159 | -99.9255
24.3196 | -99.9272
24.3233 | -99.9292
24.3271 | -99.9313
24.3308 | -99.9325
24.3345 | -99.9343
24.3382 | -99.9351
24.3419 | -99.936
24.3456 | -99.9374
24.3493 | -99.9382
24.353 | -99.9394
24.3567 | -99.9402
24.3604 | -99.9392
24.3641 | -99.9386
24.3678 | -99.9385
24.3715 | -99.9387
24.3752 | -99.9389
24.3789 | -99.9401
24.3826 | -99.9418
24.3863 | -99.9438
24.39 | -99.9464
24.3937 | -99.949
24.3974 | -99.954
24.4011 | -99.9595
24.4048 | -99.9653
24.4085 | -99.9719
24.4122 | -99.98
24.4159 | -99.9905
24.4196 | -99.9997
24.4234 | -100.0089
24.4271 | -100.0197
24.4308 | -100.0317
24.4345 | -100.0444
24.4382 | -100.0575
24.4419 | -100.0734
24.4456 | -100.0916
24.4493 | -100.1082
24.453 | -100.1262
24.4567 | -100.147
24.4604 | -100.1689
24.4641 | -100.1942
24.4678 | -100.2195
24.4715 | -100.2443
24.4752 | -100.2702
24.4789 | -100.2949
24.4826 | -100.3202
24.4863 | -100.3461
24.49 | -100.3726
24.4937 | -100.4
24.4974 | -100.4282
24.5011 | -100.4586
24.5048 | -100.4896
24.5085 | -100.5201
24.5122 | -100.5508
24.516 | -100.5837
24.5197 | -100.6159
24.5234 | -100.65
24.5271 | -100.6838
24.5308 | -100.7158
24.5345 | -100.7483
24.5382 | -100.7811
24.5419 | -100.813
24.5456 | -100.8448
24.5493 | -100.8771
24.553 | -100.9086
24.5567 | -100.9353
24.5604 | -100.9623
24.5641 | -100.9898
24.5678 | -101.0169
24.5715 | -101.0426
24.5752 | -101.0679
24.5789 | -101.0784
24.5826 | -101.0934
24.5863 | -101.0929
24.59 | -101.0681
24.5937 | -101.0353
24.5974 | -101.0092
24.6011 | -100.99
24.6048 | -100.9736
24.6085 | -100.9534
24.6123 | -100.9277
24.616 | -100.8965
24.6197 | -100.8634
24.6234 | -100.8276
24.6271 | -100.79
24.6308 | -100.7466
24.6345 | -100.7003
24.6382 | -100.6515
24.6419 | -100.6009
24.6456 | -100.5486
24.6493 | -100.4916
24.653 | -100.4302
24.6567 | -100.3666
24.6604 | -100.2976
24.6641 | -100.2251
24.6678 | -100.1517
24.6715 | -100.0766
24.6752 | -99.9989
24.6789 | -99.9211
24.6826 | -99.8422
24.6863 | -99.7634
24.69 | -99.6859
24.6937 | -99.6088
24.6974 | -99.5318
24.7011 | -99.4551
24.7048 | -99.3799
24.7086 | -99.3062
24.7123 | -99.2346
24.716 | -99.1654
24.7197 | -99.0987
24.7234 | -99.0353
24.7271 | -98.9744
24.7308 | -98.9172
24.7345 | -98.8647
24.7382 | -98.8159
24.7419 | -98.772
24.7456 | -98.7316
24.7493 | -98.695
24.753 | -98.6637
24.7567 | -98.6373
24.7604 | -98.6151
24.7641 | -98.5977
24.7678 | -98.5854
24.7715 | -98.5793
24.7752 | -98.5779
24.7789 | -98.5809
24.7826 | -98.5893
24.7863 | -98.6029
24.79 | -98.6213
24.7937 | -98.6452
24.7974 | -98.6735
24.8011 | -98.7058
24.8049 | -98.7429
24.8086 | -98.7848
24.8123 | -98.8314
24.816 | -98.883
24.8197 | -98.9384
24.8234 | -98.997
24.8271 | -99.0602
24.8308 | -99.1269
24.8345 | -99.1957
24.8382 | -99.2678
24.8419 | -99.3435
24.8456 | -99.4224
24.8493 | -99.5033
24.853 | -99.5855
24.8567 | -99.6693
24.8604 | -99.7546
24.8641 | -99.8413
24.8678 | -99.9301
24.8715 | -100.0187
24.8752 | -100.1072
24.8789 | -100.1964
24.8826 | -100.2833
24.8863 | -100.369
24.89 | -100.4535
24.8937 | -100.5379
24.8974 | -100.6216
24.9012 | -100.7037
24.9049 | -100.7857
24.9086 | -100.8642
24.9123 | -100.9382
24.916 | -101.0086
24.9197 | -101.0785
24.9234 | -101.148
24.9271 | -101.2137
24.9308 | -101.2776
24.9345 | -101.3444
24.9382 | -101.42
24.9419 | -101.5113
24.9456 | -101.5983
24.9493 | -101.6654
24.953 | -101.7131
24.9567 | -101.7496
24.9604 | -101.7824
24.9641 | -101.8135
24.9678 | -101.843
24.9715 | -101.8691
24.9752 | -101.8893
24.9789 | -101.9068
24.9826 | -101.9242
24.9863 | -101.9398
24.99 | -101.954
24.9937 | -101.9656
24.9975 | -101.9752
25.0012 | -101.9846
25.0049 | -101.9938
25.0086 | -102.0024
25.0123 | -102.0134
25.016 | -102.0228
25.0197 | -102.0294
25.0234 | -102.033
25.0271 | -102.0372
25.0308 | -102.0419
25.0345 | -102.0459
25.0382 | -102.0464
25.0419 | -102.0461
25.0456 | -102.0467
25.0493 | -102.0462
25.053 | -102.0457
25.0567 | -102.0452
25.0604 | -102.0439
25.0641 | -102.0428
25.0678 | -102.0386
25.0715 | -102.0301
25.0752 | -102.0207
25.0789 | -102.0147
25.0826 | -102.009
25.0863 | -102.003
25.0901 | -101.9964
25.0938 | -101.9869
25.0975 | -101.9758
25.1012 | -101.9647
25.1049 | -101.9524
25.1086 | -101.9395
25.1123 | -101.9218
25.116 | -101.9096
25.1197 | -101.8991
25.1234 | -101.8891
25.1271 | -101.8765
25.1308 | -101.8628
25.1345 | -101.85
25.1382 | -101.8386
25.1419 | -101.8251
25.1456 | -101.8116
25.1493 | -101.7954
25.153 | -101.7807
25.1567 | -101.7684
25.1604 | -101.756
25.1641 | -101.7427
25.1678 | -101.7291
25.1715 | -101.7127
25.1752 | -101.6966
25.1789 | -101.6824
25.1826 | -101.6686
25.1864 | -101.656
25.1901 | -101.6458
25.1938 | -101.6324
25.1975 | -101.6168
25.2012 | -101.6039
25.2049 | -101.5904
25.2086 | -101.5806
25.2123 | -101.5703
25.216 | -101.5594
25.2197 | -101.5488
25.2234 | -101.5361
25.2271 | -101.5272
25.2308 | -101.5213
25.2345 | -101.513
25.2382 | -101.5062
25.2419 | -101.5018
25.2456 | -101.4969
25.2493 | -101.4926
25.253 | -101.4877
25.2567 | -101.4821
25.2604 | -101.4759
25.2641 | -101.4688
25.2678 | -101.4627
25.2715 | -101.4602
25.2752 | -101.4576
25.2789 | -101.4546
25.2827 | -101.4521
25.2864 | -101.4499
25.2901 | -101.4484
25.2938 | -101.4465
25.2975 | -101.4445
25.3012 | -101.4435
25.3049 | -101.4405
25.3086 | -101.439
25.3123 | -101.4397
25.316 | -101.4402
25.3197 | -101.4411
25.3234 | -101.4418
25.3271 | -101.4429
25.3308 | -101.4439
25.3345 | -101.4424
25.3382 | -101.4415
25.3419 | -101.4411
25.3456 | -101.442
25.3493 | -101.4441
25.353 | -101.4458
25.3567 | -101.4481
25.3604 | -101.45
25.3641 | -101.4512
25.3678 | -101.449
25.3715 | -101.4469
25.3752 | -101.4471
25.379 | -101.4504
25.3827 | -101.4536
25.3864 | -101.455
25.3901 | -101.4569
25.3938 | -101.46
25.3975 | -101.4627
25.4012 | -101.4632
25.4049 | -101.4647
25.4086 | -101.4662
25.4123 | -101.4668
25.416 | -101.4689
25.4197 | -101.4711
25.4234 | -101.4759
25.4271 | -101.4836
25.4308 | -101.4908
25.4345 | -101.4959
25.4382 | -101.5017
25.4419 | -101.5048
25.4456 | -101.5058
25.4493 | -101.5103
25.453 | -101.5172
25.4567 | -101.5247
25.4604 | -101.5322
25.4641 | -101.5407
25.4678 | -101.5488
25.4716 | -101.557
25.4753 | -101.5658
25.479 | -101.5729
25.4827 | -101.5796
25.4864 | -101.5871
25.4901 | -101.5958
25.4938 | -101.605
25.4975 | -101.6129
25.5012 | -101.6222
25.5049 | -101.6319
25.5086 | -101.6433
25.5123 | -101.6539
25.516 | -101.6614
25.5197 | -101.6683
25.5234 | -101.6772
25.5271 | -101.6872
25.5308 | -101.6985
25.5345 | -101.7096
25.5382 | -101.7187
25.5419 | -101.7275
25.5456 | -101.7356
25.5493 | -101.7441
25.553 | -101.7558
25.5567 | -101.7678
25.5604 | -101.7762
25.5641 | -101.7866
25.5678 | -101.7948
25.5716 | -101.803
25.5753 | -101.8105
25.579 | -101.8166
25.5827 | -101.8258
25.5864 | -101.8343
25.5901 | -101.8412
25.5938 | -101.8499
25.5975 | -101.8583
25.6012 | -101.867
25.6049 | -101.8761
25.6086 | -101.8843
25.6123 | -101.8925
25.616 | -101.9028
25.6197 | -101.9136
25.6234 | -101.9254
25.6271 | -101.9362
25.6308 | -101.9448
25.6345 | -101.9539
25.6382 | -101.9626
25.6419 | -101.9726
25.6456 | -101.9807
25.6493 | -101.9895
25.653 | -101.997
25.6567 | -102.0047
25.6604 | -102.0137
25.6642 | -102.021
25.6679 | -102.0302
25.6716 | -102.0382
25.6753 | -102.0448
25.679 | -102.0513
25.6827 | -102.0573
25.6864 | -102.0655
25.6901 | -102.0754
25.6938 | -102.086
25.6975 | -102.0969
25.7012 | -102.1066
25.7049 | -102.1173
25.7086 | -102.1272
25.7123 | -102.1357
25.716 | -102.1434
25.7197 | -102.1506
25.7234 | -102.1588
25.7271 | -102.1688
25.7308 | -102.1823
25.7345 | -102.1956
25.7382 | -102.2084
25.7419 | -102.2208
25.7456 | -102.234
25.7493 | -102.2477
25.753 | -102.2613
25.7567 | -102.2763
25.7605 | -102.2893
25.7642 | -102.3032
25.7679 | -102.3185
25.7716 | -102.3349
25.7753 | -102.3522
25.779 | -102.3689
25.7827 | -102.3833
25.7864 | -102.3969
25.7901 | -102.4126
25.7938 | -102.4298
25.7975 | -102.4457
25.8012 | -102.4594
25.8049 | -102.4764
25.8086 | -102.497
25.8123 | -102.5119
25.816 | -102.5248
25.8197 | -102.5396
25.8234 | -102.5548
25.8271 | -102.5694
25.8308 | -102.5827
25.8345 | -102.5958
25.8382 | -102.6073
25.8419 | -102.619
25.8456 | -102.6315
25.8493 | -102.6432
25.853 | -102.6534
25.8568 | -102.6642
25.8605 | -102.6752
25.8642 | -102.6863
25.8679 | -102.6959
25.8716 | -102.704
25.8753 | -102.7124
25.879 | -102.7228
25.8827 | -102.7315
25.8864 | -102.7383
25.8901 | -102.742
25.8938 | -102.7453
25.8975 | -102.7499
25.9012 | -102.756
25.9049 | -102.7611
25.9086 | -102.7646
25.9123 | -102.7655
25.916 | -102.7667
25.9197 | -102.7695
25.9234 | -102.7727
25.9271 | -102.7746
25.9308 | -102.7771
25.9345 | -102.7809
25.9382 | -102.7839
25.9419 | -102.7854
25.9456 | -102.7858
25.9493 | -102.7849
25.9531 | -102.7849
25.9568 | -102.7846
25.9605 | -102.7845
25.9642 | -102.7838
25.9679 | -102.7836
25.9716 | -102.7814
25.9753 | -102.7794
25.979 | -102.7773
25.9827 | -102.7758
25.9864 | -102.7731
25.9901 | -102.7728
25.9938 | -102.7727
25.9975 | -102.7712
26.0012 | -102.771
26.0049 | -102.7705
26.0086 | -102.7699
26.0123 | -102.7719
26.016 | -102.7728
26.0197 | -102.7723
26.0234 | -102.7715
26.0271 | -102.7699
26.0308 | -102.7693
26.0345 | -102.7694
26.0382 | -102.7719
26.0419 | -102.7737
26.0456 | -102.7754
26.0494 | -102.7748
26.0531 | -102.7723
26.0568 | -102.769
26.0605 | -102.7643
26.0642 | -102.7616
26.0679 | -102.7613
26.0716 | -102.7601
26.0753 | -102.7597
26.079 | -102.7582
26.0827 | -102.7582
26.0864 | -102.7587
26.0901 | -102.7588
26.0938 | -102.7587
26.0975 | -102.7572
26.1012 | -102.7557
26.1049 | -102.7517
26.1086 | -102.7477
26.1123 | -102.7432
26.116 | -102.7385
26.1197 | -102.7348
26.1234 | -102.7319
26.1271 | -102.7301
26.1308 | -102.7281
26.1345 | -102.7262
26.1382 | -102.7236
26.142 | -102.7202
26.1457 | -102.7167
26.1494 | -102.7145
26.1531 | -102.7119
26.1568 | -102.7092
26.1605 | -102.706
26.1642 | -102.702
26.1679 | -102.6986
26.1716 | -102.6979
26.1753 | -102.6936
26.179 | -102.6918
26.1827 | -102.6933
26.1864 | -102.6934
26.1901 | -102.6904
26.1938 | -102.6875
26.1975 | -102.6841
26.2012 | -102.682
26.2049 | -102.6802
26.2086 | -102.6796
26.2123 | -102.6782
26.216 | -102.6758
26.2197 | -102.6742
26.2234 | -102.6728
26.2271 | -102.6709
26.2308 | -102.6685
26.2345 | -102.665
26.2383 | -102.6623
26.242 | -102.6618
26.2457 | -102.6598
26.2494 | -102.659
26.2531 | -102.659
26.2568 | -102.6557
26.2605 | -102.6545
26.2642 | -102.654
26.2679 | -102.6519
26.2716 | -102.6515
26.2753 | -102.6512
26.279 | -102.6498
26.2827 | -102.649
26.2864 | -102.6468
26.2901 | -102.6446
26.2938 | -102.6438
26.2975 | -102.6425
26.3012 | -102.6416
26.3049 | -102.6414
26.3086 | -102.6418
26.3123 | -102.6427
26.316 | -102.6448
26.3197 | -102.6462
26.3234 | -102.6474
26.3271 | -102.6489
26.3308 | -102.6521
26.3346 | -102.655
26.3383 | -102.6588
26.342 | -102.6644
26.3457 | -102.6707
26.3494 | -102.6765
26.3531 | -102.6822
26.3568 | -102.6891
26.3605 | -102.6953
26.3642 | -102.7023
26.3679 | -102.71
26.3716 | -102.7168
26.3753 | -102.7239
26.379 | -102.7312
26.3827 | -102.7386
26.3864 | -102.7476
26.3901 | -102.7591
26.3938 | -102.7724
26.3975 | -102.7862
26.4012 | -102.7979
26.4049 | -102.809
26.4086 | -102.8188
26.4123 | -102.8291
26.416 | -102.8411
26.4197 | -102.8543
26.4234 | -102.8663
26.4271 | -102.8811
26.4309 | -102.8961
26.4346 | -102.907
26.4383 | -102.9166
26.442 | -102.9256
26.4457 | -102.9327
26.4494 | -102.9411
26.4531 | -102.9487
26.4568 | -102.9567
26.4605 | -102.9641
26.4642 | -102.9705
26.4679 | -102.9758
26.4716 | -102.9798
26.4753 | -102.9835
26.479 | -102.9873
26.4827 | -102.9895
26.4864 | -102.9913
26.4901 | -102.9929
26.4938 | -102.9939
26.4975 | -102.9933
26.5012 | -102.991
26.5049 | -102.9857
26.5086 | -102.981
26.5123 | -102.9748
26.516 | -102.9695
26.5197 | -102.9611
26.5235 | -102.9531
26.5272 | -102.9426
26.5309 | -102.9319
26.5346 | -102.9229
26.5383 | -102.912
26.542 | -102.8964
26.5457 | -102.8844
26.5494 | -102.8714
26.5531 | -102.8557
26.5568 | -102.8392
26.5605 | -102.8246
26.5642 | -102.8081
26.5679 | -102.7925
26.5716 | -102.7777
26.5753 | -102.7593
26.579 | -102.7429
26.5827 | -102.7258
26.5864 | -102.7101
26.5901 | -102.6943
26.5938 | -102.6786
26.5975 | -102.6645
26.6012 | -102.6509
26.6049 | -102.6349
26.6086 | -102.6212
26.6123 | -102.6084
26.616 | -102.5954
26.6198 | -102.5848
26.6235 | -102.5739
26.6272 | -102.5631
26.6309 | -102.5555
26.6346 | -102.5498
26.6383 | -102.544
26.642 | -102.5371
26.6457 | -102.5273
26.6494 | -102.5213
26.6531 | -102.5187
26.6568 | -102.5156
26.6605 | -102.5133
26.6642 | -102.5123
26.6679 | -102.5109
26.6716 | -102.5104
26.6753 | -102.5109
26.679 | -102.5124
26.6827 | -102.5133
26.6864 | -102.5145
26.6901 | -102.515
26.6938 | -102.5153
26.6975 | -102.5175
26.7012 | -102.5218
26.7049 | -102.5241
26.7086 | -102.5257
26.7123 | -102.5293
26.7161 | -102.5349
26.7198 | -102.5389
26.7235 | -102.5427
26.7272 | -102.5452
26.7309 | -102.5495
26.7346 | -102.5557
26.7383 | -102.5612
26.742 | -102.5657
26.7457 | -102.5704
26.7494 | -102.5754
26.7531 | -102.5805
26.7568 | -102.5849
26.7605 | -102.5884
26.7642 | -102.591
26.7679 | -102.5943
26.7716 | -102.5972
26.7753 | -102.5996
26.779 | -102.6048
26.7827 | -102.6112
26.7864 | -102.6175
26.7901 | -102.6231
26.7938 | -102.6284
26.7975 | -102.6334
26.8012 | -102.6387
26.8049 | -102.6421
26.8086 | -102.6462
26.8124 | -102.6523
26.8161 | -102.6571
26.8198 | -102.6613
26.8235 | -102.6642
26.8272 | -102.6653
26.8309 | -102.6651
26.8346 | -102.6685
26.8383 | -102.6721
26.842 | -102.674
26.8457 | -102.6758
26.8494 | -102.6773
26.8531 | -102.6782
26.8568 | -102.6767
26.8605 | -102.676
26.8642 | -102.6763
26.8679 | -102.6754
26.8716 | -102.6733
26.8753 | -102.672
26.879 | -102.6701
26.8827 | -102.668
26.8864 | -102.6661
26.8901 | -102.6664
26.8938 | -102.6651
26.8975 | -102.667
26.9012 | -102.6678
26.9049 | -102.6664
26.9087 | -102.6658
26.9124 | -102.6663
26.9161 | -102.6671
26.9198 | -102.666
26.9235 | -102.6626
26.9272 | -102.6599
26.9309 | -102.6561
26.9346 | -102.6528
26.9383 | -102.6499
26.942 | -102.647
26.9457 | -102.6441
26.9494 | -102.6404
26.9531 | -102.6402
26.9568 | -102.639
26.9605 | -102.6357
26.9642 | -102.632
26.9679 | -102.6306
26.9716 | -102.629
26.9753 | -102.6254
26.979 | -102.6221
26.9827 | -102.6224
26.9864 | -102.6201
26.9901 | -102.6173
26.9938 | -102.6142
26.9975 | -102.612
27.0012 | -102.6103
27.005 | -102.609
27.0087 | -102.6083
27.0124 | -102.6068
27.0161 | -102.6047
27.0198 | -102.6056
27.0235 | -102.6095
27.0272 | -102.6118
27.0309 | -102.6125
27.0346 | -102.6138
27.0383 | -102.6143
27.042 | -102.6144
27.0457 | -102.6177
27.0494 | -102.6223
27.0531 | -102.6257
27.0568 | -102.6278
27.0605 | -102.6318
27.0642 | -102.6385
27.0679 | -102.6447
27.0716 | -102.651
27.0753 | -102.6546
27.079 | -102.6605
27.0827 | -102.6677
27.0864 | -102.6749
27.0901 | -102.6803
27.0938 | -102.6855
27.0976 | -102.6938
27.1013 | -102.6994
27.105 | -102.7051
27.1087 | -102.7128
27.1124 | -102.7189
27.1161 | -102.7246
27.1198 | -102.7307
27.1235 | -102.7372
27.1272 | -102.745
27.1309 | -102.7523
27.1346 | -102.7563
27.1383 | -102.7614
27.142 | -102.7671
27.1457 | -102.7742
27.1494 | -102.7783
27.1531 | -102.7789
27.1568 | -102.7805
27.1605 | -102.7822
27.1642 | -102.784
27.1679 | -102.7865
27.1716 | -102.7893
27.1753 | -102.7901
27.179 | -102.7929
27.1827 | -102.7962
27.1864 | -102.7973
27.1901 | -102.7944
27.1939 | -102.7922
27.1976 | -102.7916
27.2013 | -102.7905
27.205 | -102.7907
27.2087 | -102.7903
27.2124 | -102.7902
27.2161 | -102.7845
27.2198 | -102.7775
27.2235 | -102.7725
27.2272 | -102.7666
27.2309 | -102.7613
27.2346 | -102.7559
27.2383 | -102.75
27.242 | -102.7436
27.2457 | -102.7357
27.2494 | -102.7252
27.2531 | -102.7171
27.2568 | -102.7089
27.2605 | -102.7007
27.2642 | -102.6929
27.2679 | -102.6853
27.2716 | -102.676
27.2753 | -102.6647
27.279 | -102.6557
27.2827 | -102.6454
27.2864 | -102.6344
27.2902 | -102.6243
27.2939 | -102.6134
27.2976 | -102.6019
27.3013 | -102.5916
27.305 | -102.5828
27.3087 | -102.5753
27.3124 | -102.5683
27.3161 | -102.5608
27.3198 | -102.5529
27.3235 | -102.5437
27.3272 | -102.5365
27.3309 | -102.5273
27.3346 | -102.5181
27.3383 | -102.5087
27.342 | -102.4992
27.3457 | -102.4919
27.3494 | -102.4837
27.3531 | -102.4765
27.3568 | -102.4709
27.3605 | -102.4657
27.3642 | -102.459
27.3679 | -102.4533
27.3716 | -102.449
27.3753 | -102.4443
27.379 | -102.4405
27.3827 | -102.4364
27.3865 | -102.4299
27.3902 | -102.4212
27.3939 | -102.414
27.3976 | -102.4088
27.4013 | -102.4057
27.405 | -102.4031
27.4087 | -102.3988
27.4124 | -102.3944
27.4161 | -102.3907
27.4198 | -102.3882
27.4235 | -102.3883
27.4272 | -102.3868
27.4309 | -102.382
27.4346 | -102.3778
27.4383 | -102.3744
27.442 | -102.3707
27.4457 | -102.3676
27.4494 | -102.3641
27.4531 | -102.3608
27.4568 | -102.3565
27.4605 | -102.3556
27.4642 | -102.3501
27.4679 | -102.3438
27.4716 | -102.3382
27.4753 | -102.3324
27.4791 | -102.3273
27.4828 | -102.3222
27.4865 | -102.3178
27.4902 | -102.3131
27.4939 | -102.3076
27.4976 | -102.3029
27.5013 | -102.2977
27.505 | -102.292
27.5087 | -102.2849
27.5124 | -102.2771
27.5161 | -102.2702
27.5198 | -102.2655
27.5235 | -102.2618
27.5272 | -102.2574
27.5309 | -102.251
27.5346 | -102.2458
27.5383 | -102.2408
27.542 | -102.2351
27.5457 | -102.2302
27.5494 | -102.2224
27.5531 | -102.214
27.5568 | -102.2066
27.5605 | -102.2003
27.5642 | -102.1921
27.5679 | -102.1835
27.5716 | -102.1789
27.5754 | -102.1769
27.5791 | -102.1739
27.5828 | -102.1694
27.5865 | -102.1653
27.5902 | -102.1617
27.5939 | -102.1585
27.5976 | -102.1538
27.6013 | -102.1486
27.605 | -102.1446
27.6087 | -102.1393
27.6124 | -102.1339
27.6161 | -102.1289
27.6198 | -102.1244
27.6235 | -102.1185
27.6272 | -102.1136
27.6309 | -102.109
27.6346 | -102.1039
27.6383 | -102.099
27.642 | -102.0954
27.6457 | -102.0919
27.6494 | -102.0874
27.6531 | -102.0827
27.6568 | -102.0781
27.6605 | -102.0717
27.6642 | -102.0668
27.6679 | -102.0609
27.6717 | -102.0553
27.6754 | -102.0496
27.6791 | -102.0433
27.6828 | -102.0351
27.6865 | -102.0263
27.6902 | -102.0178
27.6939 | -102.0106
27.6976 | -102.0036
27.7013 | -101.9955
27.705 | -101.9867
27.7087 | -101.9773
27.7124 | -101.9706
27.7161 | -101.9613
27.7198 | -101.9512
27.7235 | -101.9397
27.7272 | -101.9281
27.7309 | -101.9188
27.7346 | -101.9084
27.7383 | -101.8981
27.742 | -101.8857
27.7457 | -101.8719
27.7494 | -101.8595
27.7531 | -101.8487
27.7568 | -101.8366
27.7605 | -101.8261
27.7642 | -101.8167
27.768 | -101.8062
27.7717 | -101.7949
27.7754 | -101.7869
27.7791 | -101.7813
27.7828 | -101.7746
27.7865 | -101.7665
27.7902 | -101.7569
27.7939 | -101.7468
27.7976 | -101.7364
27.8013 | -101.7276
27.805 | -101.7204
27.8087 | -101.713
27.8124 | -101.7055
27.8161 | -101.6972
27.8198 | -101.6904
27.8235 | -101.6847
27.8272 | -101.6795
27.8309 | -101.6752
27.8346 | -101.6721
27.8383 | -101.6689
27.842 | -101.6632
27.8457 | -101.6579
27.8494 | -101.6542
27.8531 | -101.6509
27.8568 | -101.6495
27.8605 | -101.6484
27.8643 | -101.647
27.868 | -101.6461
27.8717 | -101.6448
27.8754 | -101.6431
27.8791 | -101.6414
27.8828 | -101.6411
27.8865 | -101.6413
27.8902 | -101.64
27.8939 | -101.6401
27.8976 | -101.6404
27.9013 | -101.6427
27.905 | -101.6452
27.9087 | -101.6475
27.9124 | -101.6494
27.9161 | -101.6497
27.9198 | -101.6501
27.9235 | -101.6522
27.9272 | -101.654
27.9309 | -101.6567
27.9346 | -101.6595
27.9383 | -101.6602
27.942 | -101.661
27.9457 | -101.6642
27.9494 | -101.6687
27.9531 | -101.6723
27.9568 | -101.6755
27.9606 | -101.6791
27.9643 | -101.6808
27.968 | -101.6821
27.9717 | -101.6835
27.9754 | -101.6861
27.9791 | -101.6887
27.9828 | -101.6916
27.9865 | -101.6949
27.9902 | -101.6998
27.9939 | -101.7054
27.9976 | -101.7102
28.0013 | -101.7154