## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 242 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 257
Hood | 02-92B/P | 183
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1004.5 | hPa |  | PRC PRO 01
Humidity | 56.3 | %/Hr
Internal temperature of calibration bench | 24.02 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1014 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.03 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3167 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 04/06/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 07/06/2018
Technical Manager | Anthony HUE | Approval |  | 13/06/2018
CEO | Robert GLEMOT | Approval |  | 13/06/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4643 | 57.6701
0.012209 | -36.0165 | 38.5332
0.018313 | -34.9464 | 28.3527
0.024417 | -34.4802 | 22.0849
0.030521 | -34.3639 | 18.3025
0.036626 | -34.2162 | 15.2475
0.04273 | -34.1206 | 12.8721
0.048834 | -34.0755 | 11.3296
0.054938 | -34.0093 | 10.1859
0.061043 | -33.9478 | 9.1053
0.067147 | -33.9372 | 8.2738
0.073251 | -33.9374 | 7.5115
0.079355 | -33.9557 | 6.8482
0.08546 | -33.9709 | 6.2489
0.091564 | -33.9678 | 5.8035
0.097668 | -33.9505 | 5.5552
0.10377 | -33.9284 | 5.3041
0.10988 | -33.911 | 5.0219
0.11598 | -33.9009 | 4.724
0.12209 | -33.901 | 4.4401
0.12819 | -33.9087 | 4.1634
0.13429 | -33.9177 | 3.8848
0.1404 | -33.9208 | 3.6079
0.1465 | -33.9226 | 3.3503
0.15261 | -33.9202 | 3.1277
0.15871 | -33.9139 | 2.9614
0.16482 | -33.9059 | 2.8535
0.17092 | -33.903 | 2.8115
0.17702 | -33.9006 | 2.7373
0.18313 | -33.8994 | 2.661
0.18923 | -33.899 | 2.5567
0.19534 | -33.8982 | 2.4357
0.20144 | -33.8976 | 2.2972
0.20754 | -33.8957 | 2.1733
0.21365 | -33.8954 | 2.0573
0.21975 | -33.8956 | 1.9569
0.22586 | -33.8971 | 1.8681
0.23196 | -33.8997 | 1.7944
0.23807 | -33.9038 | 1.7328
0.24417 | -33.9078 | 1.6761
0.25027 | -33.9118 | 1.6142
0.25638 | -33.9142 | 1.5526
0.26248 | -33.916 | 1.4863
0.26859 | -33.9169 | 1.4198
0.27469 | -33.9172 | 1.3454
0.2808 | -33.9182 | 1.2701
0.2869 | -33.9191 | 1.1978
0.293 | -33.92 | 1.1313
0.29911 | -33.9211 | 1.0722
0.30521 | -33.9222 | 1.017
0.31132 | -33.9228 | 0.98455
0.31742 | -33.9238 | 0.94635
0.32353 | -33.9238 | 0.91654
0.32963 | -33.9239 | 0.89306
0.33573 | -33.9242 | 0.87686
0.34184 | -33.9248 | 0.85754
0.34794 | -33.9261 | 0.8303
0.35405 | -33.9275 | 0.80384
0.36015 | -33.9295 | 0.77482
0.36626 | -33.9311 | 0.74175
0.37236 | -33.933 | 0.69871
0.37846 | -33.9352 | 0.64669
0.38457 | -33.9368 | 0.58301
0.39067 | -33.9382 | 0.5304
0.39678 | -33.9393 | 0.47431
0.40288 | -33.9403 | 0.41747
0.40899 | -33.9409 | 0.36323
0.41509 | -33.9406 | 0.31086
0.42119 | -33.9408 | 0.27109
0.4273 | -33.941 | 0.22717
0.4334 | -33.9411 | 0.18334
0.43951 | -33.9413 | 0.14389
0.44561 | -33.942 | 0.10799
0.45172 | -33.9429 | 0.079727
0.45782 | -33.9433 | 0.04648
0.46392 | -33.944 | 0.012131
0.47003 | -33.9446 | -0.018346
0.47613 | -33.9447 | -0.050465
0.48224 | -33.9446 | -0.084074
0.48834 | -33.9431 | -0.11854
0.49445 | -33.9423 | -0.1526
0.50055 | -33.9412 | -0.18414
0.50665 | -33.9401 | -0.21855
0.51276 | -33.9387 | -0.25201
0.51886 | -33.9376 | -0.28594
0.52497 | -33.9368 | -0.30957
0.53107 | -33.9361 | -0.33942
0.53717 | -33.9358 | -0.36712
0.54328 | -33.9355 | -0.38976
0.54938 | -33.936 | -0.41139
0.55549 | -33.9367 | -0.43151
0.56159 | -33.9389 | -0.44207
0.5677 | -33.9401 | -0.45604
0.5738 | -33.942 | -0.47204
0.5799 | -33.9437 | -0.48669
0.58601 | -33.9458 | -0.49672
0.59211 | -33.948 | -0.50256
0.59822 | -33.95 | -0.51186
0.60432 | -33.9522 | -0.52354
0.61043 | -33.954 | -0.53156
0.61653 | -33.9557 | -0.54231
0.62263 | -33.9576 | -0.55218
0.62874 | -33.9581 | -0.56395
0.63484 | -33.959 | -0.5743
0.64095 | -33.9595 | -0.58562
0.64705 | -33.9597 | -0.59636
0.65316 | -33.9597 | -0.60984
0.65926 | -33.9595 | -0.62651
0.66536 | -33.9579 | -0.64529
0.67147 | -33.9572 | -0.65935
0.67757 | -33.9563 | -0.67384
0.68368 | -33.9551 | -0.69372
0.68978 | -33.9539 | -0.7121
0.69589 | -33.9523 | -0.73179
0.70199 | -33.9508 | -0.75473
0.70809 | -33.9494 | -0.77509
0.7142 | -33.9481 | -0.79594
0.7203 | -33.9463 | -0.81609
0.72641 | -33.9451 | -0.83846
0.73251 | -33.9439 | -0.85838
0.73862 | -33.9436 | -0.88085
0.74472 | -33.9427 | -0.90157
0.75082 | -33.9419 | -0.92372
0.75693 | -33.9416 | -0.94649
0.76303 | -33.9417 | -0.96865
0.76914 | -33.9432 | -0.98985
0.77524 | -33.9438 | -1.0095
0.78135 | -33.9446 | -1.0304
0.78745 | -33.9456 | -1.051
0.79355 | -33.9465 | -1.0712
0.79966 | -33.9481 | -1.0914
0.80576 | -33.9503 | -1.1103
0.81187 | -33.9517 | -1.1323
0.81797 | -33.9536 | -1.1532
0.82408 | -33.9551 | -1.1733
0.83018 | -33.9569 | -1.1934
0.83628 | -33.9587 | -1.2113
0.84239 | -33.9601 | -1.2302
0.84849 | -33.9617 | -1.2506
0.8546 | -33.9633 | -1.2696
0.8607 | -33.9648 | -1.2889
0.86681 | -33.9661 | -1.3068
0.87291 | -33.9673 | -1.3247
0.87901 | -33.9681 | -1.342
0.88512 | -33.9691 | -1.3592
0.89122 | -33.9702 | -1.3766
0.89733 | -33.971 | -1.3917
0.90343 | -33.9718 | -1.4089
0.90953 | -33.9724 | -1.4233
0.91564 | -33.9728 | -1.4392
0.92174 | -33.9735 | -1.4535
0.92785 | -33.9739 | -1.4683
0.93395 | -33.9745 | -1.4829
0.94006 | -33.975 | -1.4962
0.94616 | -33.9751 | -1.5097
0.95226 | -33.9757 | -1.5233
0.95837 | -33.9761 | -1.5378
0.96447 | -33.9765 | -1.5519
0.97058 | -33.9767 | -1.5665
0.97668 | -33.9769 | -1.5809
0.98279 | -33.9769 | -1.5956
0.98889 | -33.977 | -1.6078
0.99499 | -33.9772 | -1.6231
1.0011 | -33.9773 | -1.6355
1.0072 | -33.9771 | -1.6484
1.0133 | -33.977 | -1.6607
1.0194 | -33.9766 | -1.6754
1.0255 | -33.9764 | -1.6906
1.0316 | -33.976 | -1.7045
1.0377 | -33.9757 | -1.7197
1.0438 | -33.9753 | -1.7347
1.0499 | -33.9744 | -1.7516
1.056 | -33.9737 | -1.7678
1.0621 | -33.9731 | -1.7838
1.0682 | -33.9727 | -1.7975
1.0743 | -33.972 | -1.814
1.0805 | -33.9715 | -1.83
1.0866 | -33.9705 | -1.8458
1.0927 | -33.9695 | -1.8612
1.0988 | -33.9689 | -1.877
1.1049 | -33.968 | -1.8921
1.111 | -33.9671 | -1.9082
1.1171 | -33.9662 | -1.9235
1.1232 | -33.9655 | -1.9386
1.1293 | -33.965 | -1.9532
1.1354 | -33.9646 | -1.9686
1.1415 | -33.9641 | -1.9842
1.1476 | -33.9633 | -1.9993
1.1537 | -33.9626 | -2.0142
1.1598 | -33.9626 | -2.0281
1.1659 | -33.9621 | -2.0411
1.172 | -33.9617 | -2.0546
1.1781 | -33.9612 | -2.0678
1.1842 | -33.9608 | -2.0816
1.1903 | -33.961 | -2.0918
1.1964 | -33.9608 | -2.104
1.2025 | -33.9607 | -2.1172
1.2086 | -33.9609 | -2.1288
1.2147 | -33.9607 | -2.1404
1.2209 | -33.9608 | -2.1506
1.227 | -33.9615 | -2.1598
1.2331 | -33.9617 | -2.1709
1.2392 | -33.962 | -2.1818
1.2453 | -33.9623 | -2.1917
1.2514 | -33.9626 | -2.2003
1.2575 | -33.9631 | -2.2096
1.2636 | -33.9638 | -2.2171
1.2697 | -33.9644 | -2.225
1.2758 | -33.9649 | -2.2349
1.2819 | -33.9656 | -2.2436
1.288 | -33.9662 | -2.2531
1.2941 | -33.967 | -2.2601
1.3002 | -33.9677 | -2.2667
1.3063 | -33.9682 | -2.2745
1.3124 | -33.9688 | -2.2826
1.3185 | -33.9695 | -2.2902
1.3246 | -33.9701 | -2.2987
1.3307 | -33.971 | -2.3062
1.3368 | -33.9719 | -2.3141
1.3429 | -33.9725 | -2.3217
1.349 | -33.973 | -2.3302
1.3551 | -33.9737 | -2.3385
1.3613 | -33.9743 | -2.3459
1.3674 | -33.9749 | -2.3555
1.3735 | -33.9753 | -2.3647
1.3796 | -33.9758 | -2.3736
1.3857 | -33.9762 | -2.3822
1.3918 | -33.9768 | -2.3925
1.3979 | -33.9773 | -2.4018
1.404 | -33.9778 | -2.414
1.4101 | -33.9782 | -2.4242
1.4162 | -33.9786 | -2.4344
1.4223 | -33.9791 | -2.4456
1.4284 | -33.9795 | -2.4555
1.4345 | -33.9798 | -2.4674
1.4406 | -33.9801 | -2.4817
1.4467 | -33.9805 | -2.4935
1.4528 | -33.9809 | -2.5056
1.4589 | -33.9814 | -2.5182
1.465 | -33.9817 | -2.5297
1.4711 | -33.9821 | -2.5418
1.4772 | -33.9826 | -2.556
1.4833 | -33.983 | -2.5693
1.4894 | -33.9836 | -2.5822
1.4955 | -33.9841 | -2.5944
1.5016 | -33.9846 | -2.608
1.5078 | -33.985 | -2.6228
1.5139 | -33.9854 | -2.6366
1.52 | -33.9858 | -2.6501
1.5261 | -33.9861 | -2.6633
1.5322 | -33.9865 | -2.6783
1.5383 | -33.9869 | -2.6919
1.5444 | -33.9874 | -2.7053
1.5505 | -33.9879 | -2.7192
1.5566 | -33.9884 | -2.7327
1.5627 | -33.9889 | -2.7475
1.5688 | -33.9895 | -2.7621
1.5749 | -33.9898 | -2.7752
1.581 | -33.9901 | -2.7887
1.5871 | -33.9905 | -2.8031
1.5932 | -33.9909 | -2.8163
1.5993 | -33.9914 | -2.8309
1.6054 | -33.9918 | -2.8445
1.6115 | -33.9923 | -2.8592
1.6176 | -33.9926 | -2.8719
1.6237 | -33.9929 | -2.8854
1.6298 | -33.9931 | -2.898
1.6359 | -33.9935 | -2.9109
1.642 | -33.994 | -2.9231
1.6482 | -33.9942 | -2.9352
1.6543 | -33.9946 | -2.9486
1.6604 | -33.995 | -2.9606
1.6665 | -33.9951 | -2.9729
1.6726 | -33.9954 | -2.9852
1.6787 | -33.9958 | -2.9973
1.6848 | -33.9957 | -3.0088
1.6909 | -33.996 | -3.0209
1.697 | -33.9963 | -3.0324
1.7031 | -33.9963 | -3.0445
1.7092 | -33.9963 | -3.0558
1.7153 | -33.9966 | -3.0671
1.7214 | -33.9964 | -3.0779
1.7275 | -33.9964 | -3.0895
1.7336 | -33.9964 | -3.1018
1.7397 | -33.9965 | -3.1131
1.7458 | -33.9965 | -3.1242
1.7519 | -33.9965 | -3.1351
1.758 | -33.9966 | -3.1451
1.7641 | -33.9968 | -3.1557
1.7702 | -33.9968 | -3.1674
1.7763 | -33.9968 | -3.1776
1.7824 | -33.9967 | -3.1885
1.7885 | -33.9967 | -3.1992
1.7947 | -33.9969 | -3.21
1.8008 | -33.9969 | -3.2202
1.8069 | -33.997 | -3.2303
1.813 | -33.9972 | -3.2412
1.8191 | -33.9972 | -3.2527
1.8252 | -33.9973 | -3.2646
1.8313 | -33.9974 | -3.2754
1.8374 | -33.9976 | -3.2862
1.8435 | -33.9977 | -3.297
1.8496 | -33.9978 | -3.309
1.8557 | -33.998 | -3.3206
1.8618 | -33.9982 | -3.3322
1.8679 | -33.9983 | -3.3434
1.874 | -33.9984 | -3.3546
1.8801 | -33.9986 | -3.3661
1.8862 | -33.9989 | -3.3771
1.8923 | -33.999 | -3.388
1.8984 | -33.9992 | -3.4002
1.9045 | -33.9993 | -3.4111
1.9106 | -33.9993 | -3.4228
1.9167 | -33.9994 | -3.4345
1.9228 | -33.9996 | -3.4461
1.9289 | -33.9997 | -3.4584
1.9351 | -33.9999 | -3.4689
1.9412 | -34.0002 | -3.4798
1.9473 | -34.0004 | -3.4918
1.9534 | -34.0007 | -3.5032
1.9595 | -34.0009 | -3.5149
1.9656 | -34.0012 | -3.5279
1.9717 | -34.0014 | -3.5395
1.9778 | -34.0017 | -3.5506
1.9839 | -34.002 | -3.5624
1.99 | -34.0021 | -3.5735
1.9961 | -34.0022 | -3.5858
2.0022 | -34.0025 | -3.5982
2.0083 | -34.0025 | -3.6096
2.0144 | -34.0027 | -3.6215
2.0205 | -34.003 | -3.634
2.0266 | -34.0034 | -3.6459
2.0327 | -34.0036 | -3.6583
2.0388 | -34.004 | -3.6705
2.0449 | -34.0042 | -3.6824
2.051 | -34.0044 | -3.6943
2.0571 | -34.0047 | -3.7063
2.0632 | -34.0049 | -3.7181
2.0693 | -34.0051 | -3.7313
2.0754 | -34.0053 | -3.7436
2.0816 | -34.0055 | -3.756
2.0877 | -34.0056 | -3.7678
2.0938 | -34.006 | -3.7802
2.0999 | -34.0061 | -3.7924
2.106 | -34.0062 | -3.8047
2.1121 | -34.0064 | -3.8172
2.1182 | -34.0066 | -3.8287
2.1243 | -34.0068 | -3.84
2.1304 | -34.0069 | -3.8519
2.1365 | -34.0072 | -3.8644
2.1426 | -34.0074 | -3.8767
2.1487 | -34.0076 | -3.888
2.1548 | -34.0079 | -3.9002
2.1609 | -34.0083 | -3.9121
2.167 | -34.0086 | -3.9236
2.1731 | -34.0088 | -3.936
2.1792 | -34.0089 | -3.9472
2.1853 | -34.0092 | -3.9595
2.1914 | -34.0094 | -3.9711
2.1975 | -34.0097 | -3.9829
2.2036 | -34.0098 | -3.9952
2.2097 | -34.0099 | -4.007
2.2158 | -34.0102 | -4.0176
2.222 | -34.0103 | -4.0285
2.2281 | -34.0105 | -4.04
2.2342 | -34.0107 | -4.0513
2.2403 | -34.0108 | -4.0619
2.2464 | -34.011 | -4.0724
2.2525 | -34.0112 | -4.0841
2.2586 | -34.0114 | -4.0949
2.2647 | -34.0116 | -4.1054
2.2708 | -34.0117 | -4.1165
2.2769 | -34.012 | -4.127
2.283 | -34.012 | -4.1379
2.2891 | -34.0122 | -4.1487
2.2952 | -34.0122 | -4.1589
2.3013 | -34.0123 | -4.1697
2.3074 | -34.0125 | -4.1798
2.3135 | -34.0127 | -4.1912
2.3196 | -34.0128 | -4.2008
2.3257 | -34.0129 | -4.2114
2.3318 | -34.013 | -4.2223
2.3379 | -34.0131 | -4.2329
2.344 | -34.0131 | -4.244
2.3501 | -34.0132 | -4.2544
2.3562 | -34.0133 | -4.2641
2.3623 | -34.0134 | -4.2742
2.3685 | -34.0134 | -4.2844
2.3746 | -34.0135 | -4.2952
2.3807 | -34.0136 | -4.3052
2.3868 | -34.0138 | -4.3154
2.3929 | -34.0139 | -4.3258
2.399 | -34.0139 | -4.3357
2.4051 | -34.0141 | -4.3458
2.4112 | -34.0142 | -4.3558
2.4173 | -34.0144 | -4.3661
2.4234 | -34.0144 | -4.3763
2.4295 | -34.0145 | -4.3866
2.4356 | -34.0146 | -4.3968
2.4417 | -34.0147 | -4.4075
2.4478 | -34.0147 | -4.418
2.4539 | -34.0149 | -4.4278
2.46 | -34.0149 | -4.4386
2.4661 | -34.015 | -4.4491
2.4722 | -34.0152 | -4.4587
2.4783 | -34.0152 | -4.4686
2.4844 | -34.0153 | -4.4792
2.4905 | -34.0154 | -4.4898
2.4966 | -34.0155 | -4.5
2.5027 | -34.0156 | -4.51
2.5089 | -34.0157 | -4.5205
2.515 | -34.0157 | -4.5309
2.5211 | -34.0158 | -4.5411
2.5272 | -34.0159 | -4.5518
2.5333 | -34.016 | -4.5621
2.5394 | -34.0161 | -4.5723
2.5455 | -34.0163 | -4.5833
2.5516 | -34.0164 | -4.5931
2.5577 | -34.0166 | -4.6042
2.5638 | -34.0167 | -4.6156
2.5699 | -34.0168 | -4.6265
2.576 | -34.0168 | -4.6376
2.5821 | -34.0169 | -4.6483
2.5882 | -34.017 | -4.6589
2.5943 | -34.0171 | -4.6693
2.6004 | -34.0173 | -4.6799
2.6065 | -34.0175 | -4.6897
2.6126 | -34.0176 | -4.6996
2.6187 | -34.0178 | -4.7098
2.6248 | -34.0179 | -4.7202
2.6309 | -34.018 | -4.7312
2.637 | -34.0183 | -4.7426
2.6431 | -34.0185 | -4.7526
2.6492 | -34.0186 | -4.7636
2.6554 | -34.0188 | -4.7741
2.6615 | -34.019 | -4.7848
2.6676 | -34.0192 | -4.7961
2.6737 | -34.0195 | -4.8071
2.6798 | -34.0198 | -4.8178
2.6859 | -34.02 | -4.8288
2.692 | -34.0201 | -4.8397
2.6981 | -34.0204 | -4.8507
2.7042 | -34.0206 | -4.8618
2.7103 | -34.0208 | -4.8722
2.7164 | -34.0211 | -4.8829
2.7225 | -34.0214 | -4.8942
2.7286 | -34.0216 | -4.9051
2.7347 | -34.0218 | -4.9154
2.7408 | -34.0222 | -4.9256
2.7469 | -34.0224 | -4.9367
2.753 | -34.0227 | -4.9472
2.7591 | -34.0231 | -4.9582
2.7652 | -34.0233 | -4.9697
2.7713 | -34.0235 | -4.9804
2.7774 | -34.0238 | -4.9916
2.7835 | -34.024 | -5.0024
2.7896 | -34.0242 | -5.0129
2.7958 | -34.0245 | -5.0237
2.8019 | -34.0249 | -5.034
2.808 | -34.0251 | -5.0449
2.8141 | -34.0255 | -5.0562
2.8202 | -34.0258 | -5.067
2.8263 | -34.0262 | -5.078
2.8324 | -34.0266 | -5.0893
2.8385 | -34.027 | -5.1002
2.8446 | -34.0274 | -5.1111
2.8507 | -34.0278 | -5.122
2.8568 | -34.0282 | -5.1325
2.8629 | -34.0285 | -5.1432
2.869 | -34.0289 | -5.1544
2.8751 | -34.0293 | -5.1651
2.8812 | -34.0298 | -5.1767
2.8873 | -34.0301 | -5.1876
2.8934 | -34.0305 | -5.199
2.8995 | -34.0308 | -5.2101
2.9056 | -34.0311 | -5.2214
2.9117 | -34.0315 | -5.2326
2.9178 | -34.0319 | -5.2437
2.9239 | -34.0324 | -5.2554
2.93 | -34.0328 | -5.2671
2.9361 | -34.0331 | -5.2781
2.9423 | -34.0335 | -5.289
2.9484 | -34.034 | -5.3003
2.9545 | -34.0344 | -5.3116
2.9606 | -34.0348 | -5.3231
2.9667 | -34.0353 | -5.3345
2.9728 | -34.0356 | -5.3461
2.9789 | -34.0361 | -5.3575
2.985 | -34.0366 | -5.3689
2.9911 | -34.037 | -5.3806
2.9972 | -34.0375 | -5.3919
3.0033 | -34.0379 | -5.4036
3.0094 | -34.0383 | -5.4149
3.0155 | -34.0388 | -5.4261
3.0216 | -34.0392 | -5.4377
3.0277 | -34.0396 | -5.4496
3.0338 | -34.04 | -5.4612
3.0399 | -34.0405 | -5.4727
3.046 | -34.0408 | -5.4845
3.0521 | -34.0411 | -5.4959
3.0582 | -34.0415 | -5.5083
3.0643 | -34.0418 | -5.5201
3.0704 | -34.0422 | -5.5324
3.0765 | -34.0425 | -5.5437
3.0827 | -34.0429 | -5.5553
3.0888 | -34.0433 | -5.5669
3.0949 | -34.0437 | -5.5785
3.101 | -34.0441 | -5.5903
3.1071 | -34.0444 | -5.6023
3.1132 | -34.0448 | -5.6145
3.1193 | -34.0452 | -5.6263
3.1254 | -34.0456 | -5.638
3.1315 | -34.046 | -5.6502
3.1376 | -34.0463 | -5.6619
3.1437 | -34.0467 | -5.6737
3.1498 | -34.0471 | -5.686
3.1559 | -34.0475 | -5.6976
3.162 | -34.0478 | -5.71
3.1681 | -34.048 | -5.7223
3.1742 | -34.0484 | -5.7341
3.1803 | -34.0487 | -5.746
3.1864 | -34.0491 | -5.7576
3.1925 | -34.0494 | -5.7695
3.1986 | -34.0497 | -5.7817
3.2047 | -34.05 | -5.794
3.2108 | -34.0503 | -5.8054
3.2169 | -34.0506 | -5.8173
3.223 | -34.0509 | -5.8293
3.2292 | -34.0511 | -5.8413
3.2353 | -34.0513 | -5.8524
3.2414 | -34.0516 | -5.8641
3.2475 | -34.0518 | -5.8759
3.2536 | -34.0522 | -5.8876
3.2597 | -34.0524 | -5.8988
3.2658 | -34.0527 | -5.91
3.2719 | -34.0529 | -5.9211
3.278 | -34.0531 | -5.933
3.2841 | -34.0534 | -5.9444
3.2902 | -34.0536 | -5.956
3.2963 | -34.0539 | -5.9676
3.3024 | -34.054 | -5.9788
3.3085 | -34.0543 | -5.9906
3.3146 | -34.0544 | -6.0021
3.3207 | -34.0547 | -6.0138
3.3268 | -34.0549 | -6.0255
3.3329 | -34.0551 | -6.0365
3.339 | -34.0553 | -6.0475
3.3451 | -34.0555 | -6.0589
3.3512 | -34.0557 | -6.0698
3.3573 | -34.0559 | -6.0813
3.3634 | -34.0561 | -6.0926
3.3696 | -34.0562 | -6.1041
3.3757 | -34.0563 | -6.1149
3.3818 | -34.0566 | -6.1258
3.3879 | -34.0567 | -6.137
3.394 | -34.0569 | -6.1485
3.4001 | -34.0571 | -6.1594
3.4062 | -34.0572 | -6.1705
3.4123 | -34.0573 | -6.1812
3.4184 | -34.0575 | -6.1925
3.4245 | -34.0577 | -6.2035
3.4306 | -34.0577 | -6.2143
3.4367 | -34.0579 | -6.225
3.4428 | -34.0581 | -6.2355
3.4489 | -34.0582 | -6.2466
3.455 | -34.0583 | -6.2576
3.4611 | -34.0585 | -6.2689
3.4672 | -34.0586 | -6.28
3.4733 | -34.0588 | -6.2911
3.4794 | -34.0589 | -6.302
3.4855 | -34.059 | -6.3125
3.4916 | -34.0592 | -6.3229
3.4977 | -34.0593 | -6.3342
3.5038 | -34.0595 | -6.3451
3.5099 | -34.0596 | -6.3555
3.5161 | -34.0597 | -6.3658
3.5222 | -34.0598 | -6.3763
3.5283 | -34.0599 | -6.3872
3.5344 | -34.06 | -6.398
3.5405 | -34.0601 | -6.4086
3.5466 | -34.0603 | -6.4195
3.5527 | -34.0603 | -6.4302
3.5588 | -34.0604 | -6.441
3.5649 | -34.0604 | -6.4516
3.571 | -34.0605 | -6.4624
3.5771 | -34.0606 | -6.4735
3.5832 | -34.0607 | -6.4834
3.5893 | -34.061 | -6.4941
3.5954 | -34.0611 | -6.5048
3.6015 | -34.0612 | -6.5157
3.6076 | -34.0614 | -6.5262
3.6137 | -34.0615 | -6.5363
3.6198 | -34.0618 | -6.5466
3.6259 | -34.0619 | -6.5576
3.632 | -34.062 | -6.5682
3.6381 | -34.0622 | -6.5789
3.6442 | -34.0623 | -6.5895
3.6503 | -34.0625 | -6.5996
3.6565 | -34.0627 | -6.6103
3.6626 | -34.0628 | -6.6208
3.6687 | -34.063 | -6.6316
3.6748 | -34.0632 | -6.6421
3.6809 | -34.0634 | -6.6527
3.687 | -34.0635 | -6.6637
3.6931 | -34.0638 | -6.6738
3.6992 | -34.064 | -6.6843
3.7053 | -34.0641 | -6.6948
3.7114 | -34.0642 | -6.7055
3.7175 | -34.0644 | -6.7161
3.7236 | -34.0647 | -6.7261
3.7297 | -34.0649 | -6.7368
3.7358 | -34.065 | -6.7476
3.7419 | -34.0652 | -6.7583
3.748 | -34.0654 | -6.7689
3.7541 | -34.0656 | -6.7794
3.7602 | -34.0659 | -6.7897
3.7663 | -34.0661 | -6.8001
3.7724 | -34.0663 | -6.8107
3.7785 | -34.0666 | -6.821
3.7846 | -34.0667 | -6.8313
3.7907 | -34.067 | -6.8421
3.7969 | -34.0673 | -6.8527
3.803 | -34.0675 | -6.8627
3.8091 | -34.0677 | -6.8732
3.8152 | -34.068 | -6.8835
3.8213 | -34.0682 | -6.8943
3.8274 | -34.0684 | -6.9045
3.8335 | -34.0688 | -6.9151
3.8396 | -34.069 | -6.9256
3.8457 | -34.0693 | -6.9363
3.8518 | -34.0696 | -6.9466
3.8579 | -34.07 | -6.9571
3.864 | -34.0702 | -6.9678
3.8701 | -34.0705 | -6.9786
3.8762 | -34.0708 | -6.9889
3.8823 | -34.071 | -6.9996
3.8884 | -34.0713 | -7.0101
3.8945 | -34.0716 | -7.021
3.9006 | -34.0719 | -7.0319
3.9067 | -34.0722 | -7.0428
3.9128 | -34.0725 | -7.0534
3.9189 | -34.0727 | -7.0637
3.925 | -34.073 | -7.0745
3.9311 | -34.0732 | -7.0846
3.9372 | -34.0735 | -7.0953
3.9434 | -34.0738 | -7.1059
3.9495 | -34.0742 | -7.1166
3.9556 | -34.0745 | -7.127
3.9617 | -34.0748 | -7.1372
3.9678 | -34.0751 | -7.1474
3.9739 | -34.0754 | -7.158
3.98 | -34.0756 | -7.1686
3.9861 | -34.0759 | -7.1792
3.9922 | -34.0761 | -7.1897
3.9983 | -34.0764 | -7.1998
4.0044 | -34.0767 | -7.21
4.0105 | -34.077 | -7.2206
4.0166 | -34.0773 | -7.2311
4.0227 | -34.0776 | -7.2416
4.0288 | -34.0779 | -7.2527
4.0349 | -34.0781 | -7.263
4.041 | -34.0783 | -7.2735
4.0471 | -34.0786 | -7.2841
4.0532 | -34.0789 | -7.2943
4.0593 | -34.0792 | -7.3051
4.0654 | -34.0795 | -7.3154
4.0715 | -34.0797 | -7.3264
4.0776 | -34.08 | -7.3368
4.0838 | -34.0803 | -7.3478
4.0899 | -34.0806 | -7.3581
4.096 | -34.0809 | -7.3685
4.1021 | -34.0811 | -7.3789
4.1082 | -34.0814 | -7.3892
4.1143 | -34.0817 | -7.3994
4.1204 | -34.082 | -7.4096
4.1265 | -34.0822 | -7.4198
4.1326 | -34.0825 | -7.43
4.1387 | -34.0828 | -7.4404
4.1448 | -34.0831 | -7.451
4.1509 | -34.0833 | -7.4615
4.157 | -34.0836 | -7.4724
4.1631 | -34.0838 | -7.4831
4.1692 | -34.0841 | -7.4935
4.1753 | -34.0844 | -7.5039
4.1814 | -34.0846 | -7.5141
4.1875 | -34.0849 | -7.5248
4.1936 | -34.0852 | -7.5354
4.1997 | -34.0854 | -7.5455
4.2058 | -34.0856 | -7.5563
4.2119 | -34.0859 | -7.5666
4.218 | -34.0861 | -7.5768
4.2241 | -34.0863 | -7.587
4.2303 | -34.0866 | -7.5974
4.2364 | -34.0869 | -7.6078
4.2425 | -34.087 | -7.6183
4.2486 | -34.0873 | -7.629
4.2547 | -34.0874 | -7.6392
4.2608 | -34.0876 | -7.6499
4.2669 | -34.0878 | -7.6599
4.273 | -34.088 | -7.6705
4.2791 | -34.0883 | -7.6809
4.2852 | -34.0885 | -7.6915
4.2913 | -34.0886 | -7.7017
4.2974 | -34.0889 | -7.7124
4.3035 | -34.0891 | -7.7231
4.3096 | -34.0893 | -7.7336
4.3157 | -34.0895 | -7.7439
4.3218 | -34.0897 | -7.7543
4.3279 | -34.09 | -7.7648
4.334 | -34.0901 | -7.7751
4.3401 | -34.0903 | -7.785
4.3462 | -34.0906 | -7.7955
4.3523 | -34.0909 | -7.8055
4.3584 | -34.091 | -7.8155
4.3645 | -34.0912 | -7.8257
4.3707 | -34.0914 | -7.8359
4.3768 | -34.0916 | -7.8462
4.3829 | -34.0919 | -7.8569
4.389 | -34.0921 | -7.8672
4.3951 | -34.0923 | -7.8774
4.4012 | -34.0924 | -7.888
4.4073 | -34.0926 | -7.8982
4.4134 | -34.0928 | -7.9085
4.4195 | -34.0931 | -7.9187
4.4256 | -34.0934 | -7.9288
4.4317 | -34.0936 | -7.9391
4.4378 | -34.0937 | -7.949
4.4439 | -34.0939 | -7.959
4.45 | -34.094 | -7.9693
4.4561 | -34.0942 | -7.9795
4.4622 | -34.0944 | -7.9899
4.4683 | -34.0946 | -8.0001
4.4744 | -34.0948 | -8.0107
4.4805 | -34.095 | -8.0206
4.4866 | -34.0952 | -8.0309
4.4927 | -34.0954 | -8.0414
4.4988 | -34.0956 | -8.0516
4.5049 | -34.0958 | -8.0615
4.511 | -34.096 | -8.0714
4.5172 | -34.0962 | -8.0816
4.5233 | -34.0964 | -8.0923
4.5294 | -34.0966 | -8.1024
4.5355 | -34.0968 | -8.1127
4.5416 | -34.097 | -8.1228
4.5477 | -34.0972 | -8.133
4.5538 | -34.0974 | -8.1429
4.5599 | -34.0976 | -8.153
4.566 | -34.0978 | -8.1632
4.5721 | -34.098 | -8.1734
4.5782 | -34.0982 | -8.1839
4.5843 | -34.0984 | -8.1942
4.5904 | -34.0987 | -8.2044
4.5965 | -34.0989 | -8.2146
4.6026 | -34.0991 | -8.2252
4.6087 | -34.0993 | -8.2353
4.6148 | -34.0995 | -8.2459
4.6209 | -34.0997 | -8.256
4.627 | -34.1 | -8.2664
4.6331 | -34.1001 | -8.2765
4.6392 | -34.1003 | -8.2868
4.6453 | -34.1005 | -8.297
4.6514 | -34.1007 | -8.3073
4.6576 | -34.1009 | -8.3179
4.6637 | -34.1011 | -8.3278
4.6698 | -34.1013 | -8.3378
4.6759 | -34.1015 | -8.3484
4.682 | -34.1017 | -8.3585
4.6881 | -34.1018 | -8.3687
4.6942 | -34.102 | -8.3789
4.7003 | -34.1022 | -8.3892
4.7064 | -34.1024 | -8.3991
4.7125 | -34.1027 | -8.4096
4.7186 | -34.1029 | -8.4196
4.7247 | -34.1031 | -8.4299
4.7308 | -34.1033 | -8.44
4.7369 | -34.1035 | -8.4503
4.743 | -34.1038 | -8.4609
4.7491 | -34.104 | -8.4714
4.7552 | -34.1042 | -8.4817
4.7613 | -34.1043 | -8.492
4.7674 | -34.1045 | -8.5021
4.7735 | -34.1047 | -8.5128
4.7796 | -34.105 | -8.5233
4.7857 | -34.1052 | -8.5336
4.7918 | -34.1054 | -8.5438
4.7979 | -34.1057 | -8.5539
4.8041 | -34.1059 | -8.5642
4.8102 | -34.1061 | -8.5742
4.8163 | -34.1063 | -8.5843
4.8224 | -34.1066 | -8.5947
4.8285 | -34.1068 | -8.6046
4.8346 | -34.1071 | -8.6147
4.8407 | -34.1072 | -8.625
4.8468 | -34.1075 | -8.6355
4.8529 | -34.1077 | -8.6461
4.859 | -34.1079 | -8.6566
4.8651 | -34.1082 | -8.6666
4.8712 | -34.1084 | -8.6767
4.8773 | -34.1086 | -8.687
4.8834 | -34.1088 | -8.6973
4.8895 | -34.1091 | -8.7073
4.8956 | -34.1092 | -8.7176
4.9017 | -34.1094 | -8.7276
4.9078 | -34.1097 | -8.7382
4.9139 | -34.1099 | -8.748
4.92 | -34.1102 | -8.758
4.9261 | -34.1104 | -8.7679
4.9322 | -34.1107 | -8.7782
4.9383 | -34.1109 | -8.7884
4.9445 | -34.1112 | -8.7987
4.9506 | -34.1114 | -8.8088
4.9567 | -34.1116 | -8.8188
4.9628 | -34.1118 | -8.8289
4.9689 | -34.1121 | -8.8395
4.975 | -34.1123 | -8.8501
4.9811 | -34.1125 | -8.86
4.9872 | -34.1128 | -8.8702
4.9933 | -34.1131 | -8.8807
4.9994 | -34.1134 | -8.8913
5.0055 | -34.1136 | -8.9016
5.0116 | -34.1139 | -8.9123
5.0177 | -34.1142 | -8.9228
5.0238 | -34.1144 | -8.9334
5.0299 | -34.1148 | -8.944
5.036 | -34.115 | -8.9543
5.0421 | -34.1152 | -8.9645
5.0482 | -34.1155 | -8.9749
5.0543 | -34.1158 | -8.9849
5.0604 | -34.116 | -8.9951
5.0665 | -34.1163 | -9.0052
5.0726 | -34.1166 | -9.0157
5.0787 | -34.1169 | -9.0263
5.0848 | -34.1171 | -9.0366
5.091 | -34.1174 | -9.0467
5.0971 | -34.1177 | -9.0567
5.1032 | -34.118 | -9.0673
5.1093 | -34.1182 | -9.0776
5.1154 | -34.1184 | -9.0877
5.1215 | -34.1187 | -9.0978
5.1276 | -34.1189 | -9.1079
5.1337 | -34.1193 | -9.1179
5.1398 | -34.1195 | -9.1279
5.1459 | -34.1198 | -9.1385
5.152 | -34.1201 | -9.1489
5.1581 | -34.1203 | -9.1594
5.1642 | -34.1206 | -9.1698
5.1703 | -34.1209 | -9.1801
5.1764 | -34.1212 | -9.1902
5.1825 | -34.1215 | -9.2003
5.1886 | -34.1217 | -9.2103
5.1947 | -34.122 | -9.2207
5.2008 | -34.1223 | -9.2311
5.2069 | -34.1225 | -9.2417
5.213 | -34.1228 | -9.2517
5.2191 | -34.1231 | -9.2621
5.2252 | -34.1234 | -9.2724
5.2314 | -34.1236 | -9.2829
5.2375 | -34.1239 | -9.293
5.2436 | -34.1242 | -9.3034
5.2497 | -34.1245 | -9.3134
5.2558 | -34.1248 | -9.3241
5.2619 | -34.1251 | -9.3347
5.268 | -34.1254 | -9.3452
5.2741 | -34.1257 | -9.3556
5.2802 | -34.126 | -9.3663
5.2863 | -34.1263 | -9.3766
5.2924 | -34.1266 | -9.3871
5.2985 | -34.1269 | -9.3975
5.3046 | -34.1272 | -9.4079
5.3107 | -34.1274 | -9.4185
5.3168 | -34.1277 | -9.4289
5.3229 | -34.128 | -9.4393
5.329 | -34.1283 | -9.4497
5.3351 | -34.1286 | -9.4598
5.3412 | -34.1289 | -9.4701
5.3473 | -34.1291 | -9.4804
5.3534 | -34.1295 | -9.491
5.3595 | -34.1297 | -9.5014
5.3656 | -34.1301 | -9.5118
5.3717 | -34.1303 | -9.5225
5.3779 | -34.1306 | -9.5333
5.384 | -34.1309 | -9.5437
5.3901 | -34.1312 | -9.5543
5.3962 | -34.1315 | -9.5646
5.4023 | -34.1318 | -9.5752
5.4084 | -34.1322 | -9.5855
5.4145 | -34.1325 | -9.5956
5.4206 | -34.1328 | -9.606
5.4267 | -34.1331 | -9.6167
5.4328 | -34.1334 | -9.6272
5.4389 | -34.1337 | -9.6373
5.445 | -34.1339 | -9.6481
5.4511 | -34.1342 | -9.6583
5.4572 | -34.1346 | -9.6688
5.4633 | -34.1348 | -9.6792
5.4694 | -34.1351 | -9.6896
5.4755 | -34.1354 | -9.6998
5.4816 | -34.1357 | -9.7102
5.4877 | -34.136 | -9.7205
5.4938 | -34.1363 | -9.7307
5.4999 | -34.1366 | -9.741
5.506 | -34.137 | -9.7513
5.5121 | -34.1373 | -9.7616
5.5183 | -34.1375 | -9.7719
5.5244 | -34.1378 | -9.7826
5.5305 | -34.1381 | -9.7929
5.5366 | -34.1384 | -9.803
5.5427 | -34.1387 | -9.8136
5.5488 | -34.139 | -9.8241
5.5549 | -34.1393 | -9.8345
5.561 | -34.1396 | -9.8451
5.5671 | -34.1399 | -9.8554
5.5732 | -34.1402 | -9.866
5.5793 | -34.1405 | -9.8768
5.5854 | -34.1408 | -9.8871
5.5915 | -34.1411 | -9.8978
5.5976 | -34.1414 | -9.9083
5.6037 | -34.1417 | -9.9188
5.6098 | -34.142 | -9.9292
5.6159 | -34.1423 | -9.9394
5.622 | -34.1426 | -9.9499
5.6281 | -34.1429 | -9.9605
5.6342 | -34.1432 | -9.9712
5.6403 | -34.1435 | -9.9816
5.6464 | -34.1438 | -9.9922
5.6525 | -34.1441 | -10.0027
5.6586 | -34.1444 | -10.0128
5.6648 | -34.1447 | -10.0233
5.6709 | -34.1451 | -10.0339
5.677 | -34.1454 | -10.0444
5.6831 | -34.1457 | -10.0549
5.6892 | -34.1459 | -10.0655
5.6953 | -34.1462 | -10.0759
5.7014 | -34.1466 | -10.0865
5.7075 | -34.1468 | -10.0969
5.7136 | -34.1471 | -10.1073
5.7197 | -34.1474 | -10.1177
5.7258 | -34.1477 | -10.1279
5.7319 | -34.148 | -10.1384
5.738 | -34.1483 | -10.1488
5.7441 | -34.1486 | -10.1592
5.7502 | -34.149 | -10.1693
5.7563 | -34.1493 | -10.1797
5.7624 | -34.1496 | -10.1904
5.7685 | -34.1499 | -10.2008
5.7746 | -34.1502 | -10.2115
5.7807 | -34.1504 | -10.222
5.7868 | -34.1507 | -10.2326
5.7929 | -34.151 | -10.243
5.799 | -34.1513 | -10.2533
5.8052 | -34.1516 | -10.2636
5.8113 | -34.1519 | -10.2739
5.8174 | -34.1521 | -10.2842
5.8235 | -34.1524 | -10.2945
5.8296 | -34.1527 | -10.3052
5.8357 | -34.153 | -10.3155
5.8418 | -34.1533 | -10.3261
5.8479 | -34.1535 | -10.3365
5.854 | -34.1538 | -10.3471
5.8601 | -34.1542 | -10.3574
5.8662 | -34.1544 | -10.3679
5.8723 | -34.1547 | -10.3782
5.8784 | -34.155 | -10.3885
5.8845 | -34.1553 | -10.3988
5.8906 | -34.1556 | -10.4093
5.8967 | -34.1558 | -10.4195
5.9028 | -34.1562 | -10.4297
5.9089 | -34.1565 | -10.44
5.915 | -34.1568 | -10.4501
5.9211 | -34.1571 | -10.4605
5.9272 | -34.1574 | -10.4706
5.9333 | -34.1577 | -10.4809
5.9394 | -34.158 | -10.4913
5.9455 | -34.1583 | -10.5015
5.9517 | -34.1586 | -10.5116
5.9578 | -34.159 | -10.5221
5.9639 | -34.1593 | -10.5321
5.97 | -34.1596 | -10.5425
5.9761 | -34.1599 | -10.5528
5.9822 | -34.1602 | -10.5633
5.9883 | -34.1605 | -10.5736
5.9944 | -34.1608 | -10.5841
6.0005 | -34.1611 | -10.5946
6.0066 | -34.1614 | -10.6047
6.0127 | -34.1618 | -10.6148
6.0188 | -34.162 | -10.6254
6.0249 | -34.1623 | -10.6358
6.031 | -34.1626 | -10.646
6.0371 | -34.1629 | -10.6563
6.0432 | -34.1632 | -10.6665
6.0493 | -34.1635 | -10.6766
6.0554 | -34.1639 | -10.6868
6.0615 | -34.1642 | -10.6971
6.0676 | -34.1645 | -10.7076
6.0737 | -34.1648 | -10.7181
6.0798 | -34.1651 | -10.7283
6.0859 | -34.1655 | -10.7386
6.0921 | -34.1657 | -10.7486
6.0982 | -34.166 | -10.7589
6.1043 | -34.1664 | -10.7691
6.1104 | -34.1667 | -10.7796
6.1165 | -34.167 | -10.7894
6.1226 | -34.1674 | -10.7997
6.1287 | -34.1677 | -10.8099
6.1348 | -34.168 | -10.8202
6.1409 | -34.1683 | -10.8304
6.147 | -34.1686 | -10.8406
6.1531 | -34.1689 | -10.8508
6.1592 | -34.1692 | -10.8609
6.1653 | -34.1695 | -10.871
6.1714 | -34.1698 | -10.8813
6.1775 | -34.1701 | -10.8917
6.1836 | -34.1704 | -10.902
6.1897 | -34.1707 | -10.9122
6.1958 | -34.171 | -10.9227
6.2019 | -34.1714 | -10.9331
6.208 | -34.1717 | -10.9434
6.2141 | -34.172 | -10.9536
6.2202 | -34.1723 | -10.9639
6.2263 | -34.1727 | -10.974
6.2325 | -34.1729 | -10.9839
6.2386 | -34.1733 | -10.9939
6.2447 | -34.1736 | -11.0044
6.2508 | -34.1739 | -11.0146
6.2569 | -34.1743 | -11.0248
6.263 | -34.1746 | -11.035
6.2691 | -34.1749 | -11.0452
6.2752 | -34.1751 | -11.0555
6.2813 | -34.1754 | -11.0657
6.2874 | -34.1757 | -11.0757
6.2935 | -34.1761 | -11.0859
6.2996 | -34.1763 | -11.0961
6.3057 | -34.1767 | -11.1064
6.3118 | -34.1769 | -11.1167
6.3179 | -34.1772 | -11.1268
6.324 | -34.1776 | -11.137
6.3301 | -34.1779 | -11.1472
6.3362 | -34.1782 | -11.1573
6.3423 | -34.1785 | -11.1676
6.3484 | -34.1789 | -11.1779
6.3545 | -34.1791 | -11.1878
6.3606 | -34.1794 | -11.1977
6.3667 | -34.1797 | -11.2077
6.3728 | -34.18 | -11.218
6.379 | -34.1804 | -11.2281
6.3851 | -34.1807 | -11.238
6.3912 | -34.1809 | -11.2481
6.3973 | -34.1812 | -11.2584
6.4034 | -34.1815 | -11.2683
6.4095 | -34.1819 | -11.2784
6.4156 | -34.1822 | -11.2886
6.4217 | -34.1825 | -11.2987
6.4278 | -34.1828 | -11.309
6.4339 | -34.1831 | -11.319
6.44 | -34.1835 | -11.3291
6.4461 | -34.1837 | -11.3395
6.4522 | -34.184 | -11.3493
6.4583 | -34.1843 | -11.3596
6.4644 | -34.1846 | -11.3699
6.4705 | -34.1849 | -11.3797
6.4766 | -34.1852 | -11.3898
6.4827 | -34.1855 | -11.4
6.4888 | -34.1858 | -11.41
6.4949 | -34.1861 | -11.4198
6.501 | -34.1864 | -11.43
6.5071 | -34.1868 | -11.4399
6.5132 | -34.1871 | -11.4501
6.5194 | -34.1874 | -11.4603
6.5255 | -34.1877 | -11.4701
6.5316 | -34.188 | -11.4804
6.5377 | -34.1883 | -11.4905
6.5438 | -34.1885 | -11.5004
6.5499 | -34.1889 | -11.5103
6.556 | -34.1892 | -11.5204
6.5621 | -34.1895 | -11.5307
6.5682 | -34.1898 | -11.5408
6.5743 | -34.1901 | -11.5509
6.5804 | -34.1904 | -11.5611
6.5865 | -34.1907 | -11.5711
6.5926 | -34.1911 | -11.5811
6.5987 | -34.1914 | -11.5909
6.6048 | -34.1917 | -11.601
6.6109 | -34.192 | -11.6111
6.617 | -34.1923 | -11.6213
6.6231 | -34.1926 | -11.6313
6.6292 | -34.1929 | -11.6414
6.6353 | -34.1932 | -11.6517
6.6414 | -34.1935 | -11.6617
6.6475 | -34.1938 | -11.6717
6.6536 | -34.1941 | -11.6817
6.6597 | -34.1944 | -11.6915
6.6659 | -34.1947 | -11.7015
6.672 | -34.1951 | -11.7115
6.6781 | -34.1954 | -11.7216
6.6842 | -34.1957 | -11.7315
6.6903 | -34.1961 | -11.7415
6.6964 | -34.1964 | -11.7515
6.7025 | -34.1967 | -11.7614
6.7086 | -34.1969 | -11.7714
6.7147 | -34.1973 | -11.7815
6.7208 | -34.1975 | -11.7916
6.7269 | -34.1978 | -11.8016
6.733 | -34.1981 | -11.8114
6.7391 | -34.1985 | -11.8214
6.7452 | -34.1988 | -11.8312
6.7513 | -34.1991 | -11.8409
6.7574 | -34.1994 | -11.851
6.7635 | -34.1997 | -11.8608
6.7696 | -34.2 | -11.8707
6.7757 | -34.2004 | -11.8807
6.7818 | -34.2007 | -11.8907
6.7879 | -34.201 | -11.9005
6.794 | -34.2013 | -11.9103
6.8001 | -34.2016 | -11.9202
6.8063 | -34.2019 | -11.9302
6.8124 | -34.2022 | -11.9403
6.8185 | -34.2025 | -11.9504
6.8246 | -34.2028 | -11.9602
6.8307 | -34.2031 | -11.9701
6.8368 | -34.2034 | -11.9803
6.8429 | -34.2037 | -11.9903
6.849 | -34.2041 | -12.0003
6.8551 | -34.2044 | -12.0102
6.8612 | -34.2047 | -12.0201
6.8673 | -34.205 | -12.0299
6.8734 | -34.2053 | -12.0399
6.8795 | -34.2056 | -12.0499
6.8856 | -34.206 | -12.0599
6.8917 | -34.2063 | -12.0697
6.8978 | -34.2066 | -12.0794
6.9039 | -34.2069 | -12.0891
6.91 | -34.2072 | -12.099
6.9161 | -34.2076 | -12.1087
6.9222 | -34.2078 | -12.1186
6.9283 | -34.2082 | -12.1282
6.9344 | -34.2085 | -12.1382
6.9405 | -34.2089 | -12.1481
6.9466 | -34.2092 | -12.1582
6.9528 | -34.2095 | -12.1683
6.9589 | -34.2099 | -12.1783
6.965 | -34.2102 | -12.1883
6.9711 | -34.2105 | -12.1984
6.9772 | -34.2109 | -12.2083
6.9833 | -34.2112 | -12.2182
6.9894 | -34.2115 | -12.2283
6.9955 | -34.2118 | -12.238
7.0016 | -34.2122 | -12.2479
7.0077 | -34.2125 | -12.2578
7.0138 | -34.2128 | -12.2676
7.0199 | -34.2131 | -12.2775
7.026 | -34.2134 | -12.2873
7.0321 | -34.2138 | -12.2969
7.0382 | -34.2141 | -12.3065
7.0443 | -34.2145 | -12.3165
7.0504 | -34.2148 | -12.3264
7.0565 | -34.2151 | -12.3363
7.0626 | -34.2154 | -12.3459
7.0687 | -34.2157 | -12.3553
7.0748 | -34.2161 | -12.3649
7.0809 | -34.2164 | -12.3747
7.087 | -34.2168 | -12.3846
7.0932 | -34.217 | -12.3944
7.0993 | -34.2174 | -12.4043
7.1054 | -34.2177 | -12.414
7.1115 | -34.2181 | -12.4236
7.1176 | -34.2184 | -12.4334
7.1237 | -34.2187 | -12.4432
7.1298 | -34.219 | -12.4529
7.1359 | -34.2193 | -12.4628
7.142 | -34.2196 | -12.4723
7.1481 | -34.22 | -12.482
7.1542 | -34.2203 | -12.4919
7.1603 | -34.2206 | -12.5016
7.1664 | -34.221 | -12.5113
7.1725 | -34.2214 | -12.5212
7.1786 | -34.2217 | -12.531
7.1847 | -34.222 | -12.5408
7.1908 | -34.2223 | -12.5508
7.1969 | -34.2226 | -12.5607
7.203 | -34.223 | -12.5703
7.2091 | -34.2233 | -12.58
7.2152 | -34.2237 | -12.5899
7.2213 | -34.224 | -12.5995
7.2274 | -34.2243 | -12.6092
7.2335 | -34.2247 | -12.6189
7.2397 | -34.225 | -12.6287
7.2458 | -34.2254 | -12.6382
7.2519 | -34.2257 | -12.648
7.258 | -34.226 | -12.6579
7.2641 | -34.2264 | -12.6677
7.2702 | -34.2267 | -12.6776
7.2763 | -34.227 | -12.6874
7.2824 | -34.2274 | -12.6973
7.2885 | -34.2277 | -12.7071
7.2946 | -34.228 | -12.7168
7.3007 | -34.2283 | -12.7266
7.3068 | -34.2287 | -12.7367
7.3129 | -34.229 | -12.7467
7.319 | -34.2293 | -12.7565
7.3251 | -34.2297 | -12.766
7.3312 | -34.23 | -12.7758
7.3373 | -34.2304 | -12.7854
7.3434 | -34.2307 | -12.795
7.3495 | -34.2311 | -12.8046
7.3556 | -34.2314 | -12.8144
7.3617 | -34.2318 | -12.8244
7.3678 | -34.2321 | -12.8341
7.3739 | -34.2324 | -12.8441
7.3801 | -34.2327 | -12.8541
7.3862 | -34.2331 | -12.8638
7.3923 | -34.2334 | -12.8737
7.3984 | -34.2338 | -12.8833
7.4045 | -34.2341 | -12.893
7.4106 | -34.2345 | -12.9027
7.4167 | -34.2348 | -12.9125
7.4228 | -34.2352 | -12.9223
7.4289 | -34.2355 | -12.9321
7.435 | -34.2359 | -12.9417
7.4411 | -34.2362 | -12.9516
7.4472 | -34.2365 | -12.9615
7.4533 | -34.2369 | -12.9711
7.4594 | -34.2372 | -12.9811
7.4655 | -34.2376 | -12.9907
7.4716 | -34.2379 | -13.0007
7.4777 | -34.2382 | -13.0105
7.4838 | -34.2386 | -13.0201
7.4899 | -34.2389 | -13.0299
7.496 | -34.2392 | -13.0395
7.5021 | -34.2396 | -13.0492
7.5082 | -34.2399 | -13.0589
7.5143 | -34.2403 | -13.0687
7.5204 | -34.2406 | -13.0787
7.5266 | -34.2409 | -13.0883
7.5327 | -34.2413 | -13.098
7.5388 | -34.2416 | -13.1081
7.5449 | -34.242 | -13.118
7.551 | -34.2423 | -13.1276
7.5571 | -34.2427 | -13.1374
7.5632 | -34.2431 | -13.1472
7.5693 | -34.2434 | -13.1569
7.5754 | -34.2438 | -13.1665
7.5815 | -34.2441 | -13.1763
7.5876 | -34.2444 | -13.1863
7.5937 | -34.2448 | -13.1961
7.5998 | -34.2452 | -13.2058
7.6059 | -34.2456 | -13.2158
7.612 | -34.2459 | -13.2256
7.6181 | -34.2462 | -13.2353
7.6242 | -34.2466 | -13.2453
7.6303 | -34.247 | -13.2554
7.6364 | -34.2473 | -13.2652
7.6425 | -34.2476 | -13.275
7.6486 | -34.2479 | -13.2848
7.6547 | -34.2483 | -13.2945
7.6608 | -34.2487 | -13.3042
7.667 | -34.249 | -13.3141
7.6731 | -34.2493 | -13.324
7.6792 | -34.2497 | -13.334
7.6853 | -34.25 | -13.3439
7.6914 | -34.2504 | -13.3539
7.6975 | -34.2507 | -13.3637
7.7036 | -34.2511 | -13.3737
7.7097 | -34.2514 | -13.3837
7.7158 | -34.2518 | -13.3935
7.7219 | -34.2522 | -13.4032
7.728 | -34.2525 | -13.4132
7.7341 | -34.2528 | -13.4233
7.7402 | -34.2532 | -13.4331
7.7463 | -34.2535 | -13.443
7.7524 | -34.2539 | -13.4531
7.7585 | -34.2542 | -13.463
7.7646 | -34.2546 | -13.4729
7.7707 | -34.2549 | -13.483
7.7768 | -34.2552 | -13.4929
7.7829 | -34.2556 | -13.5031
7.789 | -34.2559 | -13.5132
7.7951 | -34.2562 | -13.523
7.8012 | -34.2566 | -13.5329
7.8073 | -34.257 | -13.543
7.8135 | -34.2573 | -13.5529
7.8196 | -34.2577 | -13.563
7.8257 | -34.2581 | -13.5731
7.8318 | -34.2584 | -13.5831
7.8379 | -34.2587 | -13.5929
7.844 | -34.2591 | -13.6028
7.8501 | -34.2594 | -13.6127
7.8562 | -34.2598 | -13.6228
7.8623 | -34.2602 | -13.6327
7.8684 | -34.2605 | -13.6427
7.8745 | -34.2608 | -13.6527
7.8806 | -34.2612 | -13.6628
7.8867 | -34.2616 | -13.6724
7.8928 | -34.2619 | -13.6821
7.8989 | -34.2623 | -13.6922
7.905 | -34.2626 | -13.7022
7.9111 | -34.2629 | -13.7121
7.9172 | -34.2633 | -13.7221
7.9233 | -34.2636 | -13.7321
7.9294 | -34.264 | -13.7421
7.9355 | -34.2643 | -13.7519
7.9416 | -34.2646 | -13.7619
7.9477 | -34.265 | -13.7718
7.9539 | -34.2653 | -13.7817
7.96 | -34.2657 | -13.7917
7.9661 | -34.2661 | -13.8017
7.9722 | -34.2664 | -13.8115
7.9783 | -34.2668 | -13.8217
7.9844 | -34.2672 | -13.8316
7.9905 | -34.2675 | -13.8417
7.9966 | -34.2678 | -13.8518
8.0027 | -34.2682 | -13.8618
8.0088 | -34.2685 | -13.8717
8.0149 | -34.2688 | -13.8818
8.021 | -34.2692 | -13.8919
8.0271 | -34.2696 | -13.9022
8.0332 | -34.2699 | -13.9121
8.0393 | -34.2703 | -13.9222
8.0454 | -34.2706 | -13.9324
8.0515 | -34.271 | -13.9423
8.0576 | -34.2714 | -13.9526
8.0637 | -34.2717 | -13.9624
8.0698 | -34.272 | -13.9722
8.0759 | -34.2723 | -13.9821
8.082 | -34.2727 | -13.9919
8.0881 | -34.273 | -14.0021
8.0942 | -34.2734 | -14.0121
8.1004 | -34.2738 | -14.0222
8.1065 | -34.2741 | -14.0321
8.1126 | -34.2745 | -14.0421
8.1187 | -34.2748 | -14.0523
8.1248 | -34.2752 | -14.062
8.1309 | -34.2755 | -14.0719
8.137 | -34.2759 | -14.0817
8.1431 | -34.2762 | -14.0918
8.1492 | -34.2766 | -14.1018
8.1553 | -34.2769 | -14.1118
8.1614 | -34.2772 | -14.1219
8.1675 | -34.2776 | -14.1317
8.1736 | -34.2779 | -14.142
8.1797 | -34.2782 | -14.1521
8.1858 | -34.2786 | -14.1621
8.1919 | -34.2789 | -14.1723
8.198 | -34.2793 | -14.1821
8.2041 | -34.2797 | -14.1922
8.2102 | -34.28 | -14.202
8.2163 | -34.2804 | -14.2118
8.2224 | -34.2807 | -14.2218
8.2285 | -34.2811 | -14.2315
8.2346 | -34.2814 | -14.2414
8.2408 | -34.2817 | -14.2512
8.2469 | -34.2821 | -14.2614
8.253 | -34.2825 | -14.2712
8.2591 | -34.2828 | -14.2812
8.2652 | -34.2832 | -14.2913
8.2713 | -34.2836 | -14.3013
8.2774 | -34.2839 | -14.3113
8.2835 | -34.2843 | -14.3216
8.2896 | -34.2846 | -14.3315
8.2957 | -34.285 | -14.3415
8.3018 | -34.2853 | -14.3514
8.3079 | -34.2857 | -14.3611
8.314 | -34.2861 | -14.3711
8.3201 | -34.2865 | -14.381
8.3262 | -34.2868 | -14.3909
8.3323 | -34.2871 | -14.401
8.3384 | -34.2875 | -14.4112
8.3445 | -34.2878 | -14.4212
8.3506 | -34.2882 | -14.4313
8.3567 | -34.2886 | -14.4414
8.3628 | -34.2889 | -14.4513
8.3689 | -34.2893 | -14.4613
8.375 | -34.2896 | -14.4713
8.3812 | -34.29 | -14.4812
8.3873 | -34.2904 | -14.4912
8.3934 | -34.2907 | -14.5013
8.3995 | -34.291 | -14.5111
8.4056 | -34.2914 | -14.521
8.4117 | -34.2917 | -14.531
8.4178 | -34.2921 | -14.541
8.4239 | -34.2924 | -14.5508
8.43 | -34.2928 | -14.5608
8.4361 | -34.2931 | -14.5709
8.4422 | -34.2935 | -14.5808
8.4483 | -34.2938 | -14.5906
8.4544 | -34.2942 | -14.6006
8.4605 | -34.2945 | -14.6104
8.4666 | -34.2949 | -14.6206
8.4727 | -34.2952 | -14.6305
8.4788 | -34.2956 | -14.6405
8.4849 | -34.296 | -14.6505
8.491 | -34.2963 | -14.6605
8.4971 | -34.2967 | -14.6703
8.5032 | -34.2971 | -14.68
8.5093 | -34.2974 | -14.6897
8.5154 | -34.2978 | -14.6997
8.5215 | -34.2982 | -14.7095
8.5277 | -34.2985 | -14.7195
8.5338 | -34.2989 | -14.7294
8.5399 | -34.2992 | -14.7394
8.546 | -34.2996 | -14.7491
8.5521 | -34.3 | -14.759
8.5582 | -34.3003 | -14.7689
8.5643 | -34.3006 | -14.7786
8.5704 | -34.301 | -14.7882
8.5765 | -34.3014 | -14.7982
8.5826 | -34.3017 | -14.8078
8.5887 | -34.3021 | -14.8177
8.5948 | -34.3025 | -14.8277
8.6009 | -34.3029 | -14.8375
8.607 | -34.3032 | -14.8475
8.6131 | -34.3036 | -14.8575
8.6192 | -34.304 | -14.8671
8.6253 | -34.3044 | -14.8768
8.6314 | -34.3047 | -14.8867
8.6375 | -34.305 | -14.8965
8.6436 | -34.3054 | -14.9063
8.6497 | -34.3057 | -14.9161
8.6558 | -34.3061 | -14.9259
8.6619 | -34.3064 | -14.9357
8.6681 | -34.3068 | -14.9457
8.6742 | -34.3072 | -14.9555
8.6803 | -34.3075 | -14.9654
8.6864 | -34.3079 | -14.975
8.6925 | -34.3083 | -14.9847
8.6986 | -34.3086 | -14.9946
8.7047 | -34.3089 | -15.0047
8.7108 | -34.3093 | -15.0145
8.7169 | -34.3097 | -15.0243
8.723 | -34.31 | -15.034
8.7291 | -34.3104 | -15.0436
8.7352 | -34.3108 | -15.0536
8.7413 | -34.3111 | -15.0635
8.7474 | -34.3115 | -15.0732
8.7535 | -34.3119 | -15.0831
8.7596 | -34.3123 | -15.0929
8.7657 | -34.3126 | -15.1029
8.7718 | -34.313 | -15.1127
8.7779 | -34.3134 | -15.1226
8.784 | -34.3137 | -15.1324
8.7901 | -34.3141 | -15.142
8.7962 | -34.3145 | -15.1516
8.8023 | -34.3148 | -15.1613
8.8084 | -34.3152 | -15.1711
8.8146 | -34.3156 | -15.1808
8.8207 | -34.3159 | -15.1905
8.8268 | -34.3163 | -15.1999
8.8329 | -34.3167 | -15.2097
8.839 | -34.3171 | -15.2194
8.8451 | -34.3174 | -15.2291
8.8512 | -34.3178 | -15.2387
8.8573 | -34.3182 | -15.2484
8.8634 | -34.3185 | -15.2579
8.8695 | -34.3189 | -15.2679
8.8756 | -34.3192 | -15.2777
8.8817 | -34.3196 | -15.2872
8.8878 | -34.32 | -15.2971
8.8939 | -34.3204 | -15.3069
8.9 | -34.3208 | -15.3167
8.9061 | -34.3211 | -15.3265
8.9122 | -34.3215 | -15.3362
8.9183 | -34.3219 | -15.346
8.9244 | -34.3223 | -15.3558
8.9305 | -34.3227 | -15.3654
8.9366 | -34.323 | -15.3749
8.9427 | -34.3234 | -15.3848
8.9488 | -34.3238 | -15.3943
8.955 | -34.3241 | -15.4039
8.9611 | -34.3245 | -15.4138
8.9672 | -34.3249 | -15.4233
8.9733 | -34.3253 | -15.4328
8.9794 | -34.3257 | -15.4424
8.9855 | -34.3261 | -15.4522
8.9916 | -34.3264 | -15.4621
8.9977 | -34.3268 | -15.4717
9.0038 | -34.3272 | -15.4814
9.0099 | -34.3276 | -15.491
9.016 | -34.328 | -15.5007
9.0221 | -34.3284 | -15.5102
9.0282 | -34.3288 | -15.5198
9.0343 | -34.3292 | -15.5295
9.0404 | -34.3296 | -15.539
9.0465 | -34.33 | -15.5488
9.0526 | -34.3303 | -15.5583
9.0587 | -34.3308 | -15.568
9.0648 | -34.3312 | -15.5775
9.0709 | -34.3315 | -15.5872
9.077 | -34.332 | -15.5968
9.0831 | -34.3323 | -15.6066
9.0892 | -34.3327 | -15.6163
9.0953 | -34.3332 | -15.626
9.1015 | -34.3335 | -15.6357
9.1076 | -34.334 | -15.645
9.1137 | -34.3343 | -15.6547
9.1198 | -34.3347 | -15.6643
9.1259 | -34.3351 | -15.6739
9.132 | -34.3355 | -15.6837
9.1381 | -34.3359 | -15.6932
9.1442 | -34.3364 | -15.7025
9.1503 | -34.3368 | -15.7122
9.1564 | -34.3372 | -15.7218
9.1625 | -34.3376 | -15.7314
9.1686 | -34.338 | -15.741
9.1747 | -34.3384 | -15.7506
9.1808 | -34.3388 | -15.76
9.1869 | -34.3392 | -15.7697
9.193 | -34.3396 | -15.7792
9.1991 | -34.34 | -15.7889
9.2052 | -34.3404 | -15.7985
9.2113 | -34.3408 | -15.8081
9.2174 | -34.3413 | -15.8176
9.2235 | -34.3417 | -15.8272
9.2296 | -34.3421 | -15.8367
9.2357 | -34.3425 | -15.8464
9.2419 | -34.3429 | -15.856
9.248 | -34.3434 | -15.8655
9.2541 | -34.3438 | -15.875
9.2602 | -34.3442 | -15.8848
9.2663 | -34.3446 | -15.8941
9.2724 | -34.345 | -15.9037
9.2785 | -34.3454 | -15.9132
9.2846 | -34.3458 | -15.9225
9.2907 | -34.3463 | -15.9321
9.2968 | -34.3467 | -15.9418
9.3029 | -34.3471 | -15.9514
9.309 | -34.3475 | -15.9611
9.3151 | -34.3478 | -15.9705
9.3212 | -34.3483 | -15.9799
9.3273 | -34.3487 | -15.9894
9.3334 | -34.3491 | -15.9988
9.3395 | -34.3495 | -16.0083
9.3456 | -34.3499 | -16.0177
9.3517 | -34.3503 | -16.0273
9.3578 | -34.3507 | -16.0368
9.3639 | -34.3511 | -16.0462
9.37 | -34.3516 | -16.0558
9.3761 | -34.352 | -16.0655
9.3822 | -34.3524 | -16.075
9.3884 | -34.3529 | -16.0842
9.3945 | -34.3533 | -16.0937
9.4006 | -34.3537 | -16.1033
9.4067 | -34.3541 | -16.1129
9.4128 | -34.3545 | -16.1225
9.4189 | -34.3549 | -16.1322
9.425 | -34.3554 | -16.1415
9.4311 | -34.3558 | -16.1511
9.4372 | -34.3563 | -16.1605
9.4433 | -34.3567 | -16.17
9.4494 | -34.3571 | -16.1793
9.4555 | -34.3575 | -16.1889
9.4616 | -34.358 | -16.1984
9.4677 | -34.3584 | -16.2078
9.4738 | -34.3588 | -16.2173
9.4799 | -34.3592 | -16.2267
9.486 | -34.3597 | -16.2362
9.4921 | -34.3601 | -16.2457
9.4982 | -34.3605 | -16.2549
9.5043 | -34.3609 | -16.2643
9.5104 | -34.3613 | -16.2736
9.5165 | -34.3618 | -16.2831
9.5226 | -34.3622 | -16.2927
9.5288 | -34.3626 | -16.302
9.5349 | -34.363 | -16.3114
9.541 | -34.3634 | -16.3208
9.5471 | -34.3638 | -16.3303
9.5532 | -34.3643 | -16.3398
9.5593 | -34.3647 | -16.3493
9.5654 | -34.3651 | -16.3586
9.5715 | -34.3655 | -16.3681
9.5776 | -34.3659 | -16.3777
9.5837 | -34.3664 | -16.3873
9.5898 | -34.3667 | -16.3968
9.5959 | -34.3672 | -16.4062
9.602 | -34.3676 | -16.4154
9.6081 | -34.368 | -16.4249
9.6142 | -34.3685 | -16.4343
9.6203 | -34.3689 | -16.4438
9.6264 | -34.3693 | -16.4531
9.6325 | -34.3697 | -16.4626
9.6386 | -34.3701 | -16.4719
9.6447 | -34.3705 | -16.4814
9.6508 | -34.3709 | -16.4908
9.6569 | -34.3713 | -16.5002
9.663 | -34.3718 | -16.5096
9.6691 | -34.3722 | -16.5192
9.6753 | -34.3726 | -16.5286
9.6814 | -34.373 | -16.538
9.6875 | -34.3734 | -16.5471
9.6936 | -34.3739 | -16.5565
9.6997 | -34.3743 | -16.5661
9.7058 | -34.3748 | -16.5756
9.7119 | -34.3752 | -16.585
9.718 | -34.3756 | -16.5946
9.7241 | -34.376 | -16.6039
9.7302 | -34.3764 | -16.6134
9.7363 | -34.3768 | -16.6227
9.7424 | -34.3773 | -16.632
9.7485 | -34.3777 | -16.6414
9.7546 | -34.3781 | -16.6507
9.7607 | -34.3785 | -16.6604
9.7668 | -34.3789 | -16.6698
9.7729 | -34.3794 | -16.6793
9.779 | -34.3798 | -16.6884
9.7851 | -34.3802 | -16.698
9.7912 | -34.3807 | -16.7073
9.7973 | -34.3811 | -16.7169
9.8034 | -34.3815 | -16.7261
9.8095 | -34.382 | -16.7357
9.8157 | -34.3824 | -16.7452
9.8218 | -34.3828 | -16.7548
9.8279 | -34.3832 | -16.7641
9.834 | -34.3836 | -16.7736
9.8401 | -34.384 | -16.7829
9.8462 | -34.3845 | -16.7921
9.8523 | -34.3849 | -16.8015
9.8584 | -34.3853 | -16.8108
9.8645 | -34.3857 | -16.8202
9.8706 | -34.3861 | -16.8297
9.8767 | -34.3865 | -16.839
9.8828 | -34.3869 | -16.8482
9.8889 | -34.3873 | -16.8575
9.895 | -34.3877 | -16.8669
9.9011 | -34.3881 | -16.8763
9.9072 | -34.3885 | -16.8857
9.9133 | -34.3889 | -16.895
9.9194 | -34.3893 | -16.9044
9.9255 | -34.3897 | -16.9137
9.9316 | -34.3901 | -16.923
9.9377 | -34.3906 | -16.9324
9.9438 | -34.3909 | -16.9419
9.9499 | -34.3914 | -16.9514
9.956 | -34.3918 | -16.9608
9.9622 | -34.3922 | -16.9703
9.9683 | -34.3926 | -16.9798
9.9744 | -34.393 | -16.9893
9.9805 | -34.3934 | -16.9987
9.9866 | -34.3938 | -17.008
9.9927 | -34.3942 | -17.0172
9.9988 | -34.3946 | -17.0266
10.0049 | -34.395 | -17.0361
10.011 | -34.3954 | -17.0456
10.0171 | -34.3958 | -17.055
10.0232 | -34.3963 | -17.0641
10.0293 | -34.3967 | -17.0737
10.0354 | -34.3971 | -17.0831
10.0415 | -34.3975 | -17.0925
10.0476 | -34.3979 | -17.1019
10.0537 | -34.3983 | -17.1115
10.0598 | -34.3987 | -17.1211
10.0659 | -34.3991 | -17.1305
10.072 | -34.3995 | -17.1398
10.0781 | -34.3999 | -17.1493
10.0842 | -34.4004 | -17.1585
10.0903 | -34.4008 | -17.168
10.0964 | -34.4012 | -17.177
10.1026 | -34.4016 | -17.1864
10.1087 | -34.402 | -17.1957
10.1148 | -34.4024 | -17.2052
10.1209 | -34.4028 | -17.2144
10.127 | -34.4032 | -17.2235
10.1331 | -34.4036 | -17.2327
10.1392 | -34.404 | -17.2419
10.1453 | -34.4044 | -17.2512
10.1514 | -34.4048 | -17.2605
10.1575 | -34.4052 | -17.2701
10.1636 | -34.4056 | -17.2792
10.1697 | -34.406 | -17.2886
10.1758 | -34.4065 | -17.2976
10.1819 | -34.4068 | -17.3069
10.188 | -34.4072 | -17.3162
10.1941 | -34.4076 | -17.3256
10.2002 | -34.408 | -17.3349
10.2063 | -34.4084 | -17.3441
10.2124 | -34.4088 | -17.3535
10.2185 | -34.4092 | -17.3631
10.2246 | -34.4096 | -17.3724
10.2307 | -34.41 | -17.3817
10.2368 | -34.4104 | -17.3908
10.2429 | -34.4108 | -17.4002
10.2491 | -34.4112 | -17.4097
10.2552 | -34.4116 | -17.4191
10.2613 | -34.412 | -17.4283
10.2674 | -34.4124 | -17.4377
10.2735 | -34.4128 | -17.4472
10.2796 | -34.4132 | -17.4566
10.2857 | -34.4136 | -17.4662
10.2918 | -34.414 | -17.4755
10.2979 | -34.4144 | -17.4849
10.304 | -34.4148 | -17.4939
10.3101 | -34.4151 | -17.5031
10.3162 | -34.4156 | -17.5126
10.3223 | -34.416 | -17.5219
10.3284 | -34.4164 | -17.5312
10.3345 | -34.4168 | -17.5405
10.3406 | -34.4172 | -17.5499
10.3467 | -34.4175 | -17.5593
10.3528 | -34.4179 | -17.5684
10.3589 | -34.4183 | -17.5776
10.365 | -34.4187 | -17.587
10.3711 | -34.4191 | -17.5963
10.3772 | -34.4196 | -17.6056
10.3833 | -34.42 | -17.6151
10.3895 | -34.4204 | -17.6242
10.3956 | -34.4208 | -17.6336
10.4017 | -34.4212 | -17.6428
10.4078 | -34.4216 | -17.6521
10.4139 | -34.422 | -17.6613
10.42 | -34.4224 | -17.6708
10.4261 | -34.4228 | -17.6799
10.4322 | -34.4232 | -17.6892
10.4383 | -34.4235 | -17.6985
10.4444 | -34.4239 | -17.7077
10.4505 | -34.4243 | -17.7171
10.4566 | -34.4246 | -17.7266
10.4627 | -34.4251 | -17.7357
10.4688 | -34.4255 | -17.7448
10.4749 | -34.4259 | -17.754
10.481 | -34.4262 | -17.7633
10.4871 | -34.4266 | -17.7724
10.4932 | -34.427 | -17.7815
10.4993 | -34.4274 | -17.7909
10.5054 | -34.4278 | -17.8001
10.5115 | -34.4282 | -17.8095
10.5176 | -34.4286 | -17.8188
10.5237 | -34.429 | -17.8283
10.5298 | -34.4294 | -17.8376
10.536 | -34.4298 | -17.847
10.5421 | -34.4302 | -17.8563
10.5482 | -34.4305 | -17.8657
10.5543 | -34.4309 | -17.875
10.5604 | -34.4313 | -17.8842
10.5665 | -34.4317 | -17.8931
10.5726 | -34.4321 | -17.9025
10.5787 | -34.4325 | -17.9115
10.5848 | -34.4328 | -17.921
10.5909 | -34.4333 | -17.9303
10.597 | -34.4336 | -17.9395
10.6031 | -34.434 | -17.9491
10.6092 | -34.4344 | -17.9583
10.6153 | -34.4348 | -17.9677
10.6214 | -34.4352 | -17.9767
10.6275 | -34.4356 | -17.9859
10.6336 | -34.4359 | -17.995
10.6397 | -34.4363 | -18.0042
10.6458 | -34.4367 | -18.0136
10.6519 | -34.4371 | -18.0229
10.658 | -34.4375 | -18.0322
10.6641 | -34.4379 | -18.0416
10.6702 | -34.4383 | -18.0511
10.6764 | -34.4387 | -18.0604
10.6825 | -34.4391 | -18.0695
10.6886 | -34.4395 | -18.0786
10.6947 | -34.4399 | -18.0879
10.7008 | -34.4403 | -18.0971
10.7069 | -34.4407 | -18.1065
10.713 | -34.4411 | -18.1158
10.7191 | -34.4415 | -18.1251
10.7252 | -34.4418 | -18.1342
10.7313 | -34.4422 | -18.1436
10.7374 | -34.4426 | -18.1528
10.7435 | -34.443 | -18.1619
10.7496 | -34.4434 | -18.1709
10.7557 | -34.4438 | -18.1803
10.7618 | -34.4442 | -18.1892
10.7679 | -34.4446 | -18.1987
10.774 | -34.445 | -18.2079
10.7801 | -34.4455 | -18.2174
10.7862 | -34.4459 | -18.2264
10.7923 | -34.4462 | -18.2361
10.7984 | -34.4466 | -18.2453
10.8045 | -34.447 | -18.2544
10.8106 | -34.4474 | -18.2641
10.8168 | -34.4478 | -18.2734
10.8229 | -34.4482 | -18.2828
10.829 | -34.4486 | -18.2921
10.8351 | -34.449 | -18.3012
10.8412 | -34.4494 | -18.31
10.8473 | -34.4498 | -18.3193
10.8534 | -34.4503 | -18.3285
10.8595 | -34.4507 | -18.3378
10.8656 | -34.4511 | -18.347
10.8717 | -34.4515 | -18.3564
10.8778 | -34.4519 | -18.3655
10.8839 | -34.4523 | -18.3747
10.89 | -34.4527 | -18.384
10.8961 | -34.4532 | -18.3932
10.9022 | -34.4536 | -18.4025
10.9083 | -34.4539 | -18.4117
10.9144 | -34.4543 | -18.4207
10.9205 | -34.4547 | -18.43
10.9266 | -34.4551 | -18.4395
10.9327 | -34.4554 | -18.4485
10.9388 | -34.4558 | -18.4573
10.9449 | -34.4562 | -18.4664
10.951 | -34.4566 | -18.4755
10.9571 | -34.4569 | -18.4846
10.9633 | -34.4574 | -18.4938
10.9694 | -34.4578 | -18.5033
10.9755 | -34.4582 | -18.5127
10.9816 | -34.4586 | -18.5217
10.9877 | -34.459 | -18.5308
10.9938 | -34.4595 | -18.54
10.9999 | -34.4599 | -18.549
11.006 | -34.4603 | -18.5581
11.0121 | -34.4607 | -18.5673
11.0182 | -34.4611 | -18.5766
11.0243 | -34.4615 | -18.5858
11.0304 | -34.4619 | -18.595
11.0365 | -34.4623 | -18.6038
11.0426 | -34.4628 | -18.613
11.0487 | -34.4631 | -18.6225
11.0548 | -34.4635 | -18.6316
11.0609 | -34.4639 | -18.6408
11.067 | -34.4644 | -18.6501
11.0731 | -34.4648 | -18.6592
11.0792 | -34.4652 | -18.6685
11.0853 | -34.4656 | -18.6778
11.0914 | -34.466 | -18.687
11.0975 | -34.4664 | -18.6961
11.1037 | -34.4668 | -18.7052
11.1098 | -34.4672 | -18.7146
11.1159 | -34.4676 | -18.7236
11.122 | -34.468 | -18.733
11.1281 | -34.4684 | -18.742
11.1342 | -34.4688 | -18.7513
11.1403 | -34.4693 | -18.7606
11.1464 | -34.4697 | -18.7697
11.1525 | -34.4702 | -18.7786
11.1586 | -34.4706 | -18.7879
11.1647 | -34.471 | -18.797
11.1708 | -34.4714 | -18.8063
11.1769 | -34.4718 | -18.8155
11.183 | -34.4723 | -18.8243
11.1891 | -34.4727 | -18.8333
11.1952 | -34.4731 | -18.8424
11.2013 | -34.4735 | -18.8518
11.2074 | -34.4739 | -18.8609
11.2135 | -34.4743 | -18.8704
11.2196 | -34.4747 | -18.8795
11.2257 | -34.4751 | -18.8885
11.2318 | -34.4756 | -18.8975
11.2379 | -34.476 | -18.9069
11.244 | -34.4765 | -18.9163
11.2502 | -34.4769 | -18.9256
11.2563 | -34.4773 | -18.9351
11.2624 | -34.4778 | -18.9441
11.2685 | -34.4782 | -18.9528
11.2746 | -34.4786 | -18.9618
11.2807 | -34.479 | -18.9709
11.2868 | -34.4794 | -18.9802
11.2929 | -34.4798 | -18.989
11.299 | -34.4803 | -18.9983
11.3051 | -34.4807 | -19.0074
11.3112 | -34.4811 | -19.0166
11.3173 | -34.4815 | -19.0256
11.3234 | -34.4819 | -19.0342
11.3295 | -34.4824 | -19.0436
11.3356 | -34.4828 | -19.0527
11.3417 | -34.4832 | -19.0616
11.3478 | -34.4836 | -19.0709
11.3539 | -34.484 | -19.0798
11.36 | -34.4845 | -19.0889
11.3661 | -34.4849 | -19.098
11.3722 | -34.4854 | -19.1068
11.3783 | -34.4858 | -19.1159
11.3844 | -34.4862 | -19.1249
11.3906 | -34.4867 | -19.1339
11.3967 | -34.4872 | -19.143
11.4028 | -34.4876 | -19.1523
11.4089 | -34.488 | -19.1615
11.415 | -34.4884 | -19.1707
11.4211 | -34.4888 | -19.1801
11.4272 | -34.4893 | -19.1893
11.4333 | -34.4897 | -19.1981
11.4394 | -34.4901 | -19.2069
11.4455 | -34.4905 | -19.2159
11.4516 | -34.4909 | -19.225
11.4577 | -34.4914 | -19.2341
11.4638 | -34.4919 | -19.2428
11.4699 | -34.4923 | -19.2519
11.476 | -34.4927 | -19.2611
11.4821 | -34.4932 | -19.2702
11.4882 | -34.4936 | -19.2792
11.4943 | -34.494 | -19.2885
11.5004 | -34.4944 | -19.2971
11.5065 | -34.4949 | -19.3059
11.5126 | -34.4953 | -19.3148
11.5187 | -34.4957 | -19.3241
11.5248 | -34.4961 | -19.3334
11.5309 | -34.4966 | -19.3427
11.5371 | -34.497 | -19.3517
11.5432 | -34.4975 | -19.3607
11.5493 | -34.4979 | -19.3701
11.5554 | -34.4983 | -19.3795
11.5615 | -34.4988 | -19.389
11.5676 | -34.4992 | -19.398
11.5737 | -34.4997 | -19.4072
11.5798 | -34.5001 | -19.4159
11.5859 | -34.5006 | -19.4248
11.592 | -34.501 | -19.4336
11.5981 | -34.5015 | -19.4426
11.6042 | -34.5019 | -19.4516
11.6103 | -34.5023 | -19.4607
11.6164 | -34.5027 | -19.4698
11.6225 | -34.5032 | -19.4786
11.6286 | -34.5036 | -19.4875
11.6347 | -34.504 | -19.4966
11.6408 | -34.5045 | -19.5053
11.6469 | -34.5049 | -19.5143
11.653 | -34.5053 | -19.523
11.6591 | -34.5057 | -19.5321
11.6652 | -34.5062 | -19.5411
11.6713 | -34.5066 | -19.5501
11.6775 | -34.5071 | -19.559
11.6836 | -34.5075 | -19.5679
11.6897 | -34.508 | -19.5769
11.6958 | -34.5084 | -19.5859
11.7019 | -34.5089 | -19.5949
11.708 | -34.5093 | -19.6041
11.7141 | -34.5099 | -19.6129
11.7202 | -34.5103 | -19.6219
11.7263 | -34.5107 | -19.6308
11.7324 | -34.5112 | -19.6401
11.7385 | -34.5116 | -19.649
11.7446 | -34.512 | -19.6578
11.7507 | -34.5125 | -19.6667
11.7568 | -34.5129 | -19.6755
11.7629 | -34.5134 | -19.6846
11.769 | -34.5138 | -19.6936
11.7751 | -34.5143 | -19.7026
11.7812 | -34.5148 | -19.7114
11.7873 | -34.5152 | -19.7204
11.7934 | -34.5156 | -19.7294
11.7995 | -34.5161 | -19.7384
11.8056 | -34.5166 | -19.7473
11.8117 | -34.517 | -19.7564
11.8178 | -34.5175 | -19.7652
11.824 | -34.5179 | -19.7741
11.8301 | -34.5184 | -19.783
11.8362 | -34.5188 | -19.792
11.8423 | -34.5193 | -19.8009
11.8484 | -34.5197 | -19.8098
11.8545 | -34.5202 | -19.8187
11.8606 | -34.5206 | -19.8275
11.8667 | -34.5211 | -19.8365
11.8728 | -34.5215 | -19.8452
11.8789 | -34.522 | -19.8541
11.885 | -34.5224 | -19.8631
11.8911 | -34.5229 | -19.8719
11.8972 | -34.5234 | -19.8807
11.9033 | -34.5238 | -19.8894
11.9094 | -34.5243 | -19.8983
11.9155 | -34.5247 | -19.9072
11.9216 | -34.5252 | -19.9158
11.9277 | -34.5256 | -19.9247
11.9338 | -34.5261 | -19.9335
11.9399 | -34.5265 | -19.9425
11.946 | -34.527 | -19.9514
11.9521 | -34.5274 | -19.9602
11.9582 | -34.5279 | -19.9688
11.9644 | -34.5283 | -19.9776
11.9705 | -34.5288 | -19.9865
11.9766 | -34.5293 | -19.9954
11.9827 | -34.5298 | -20.0042
11.9888 | -34.5302 | -20.0128
11.9949 | -34.5307 | -20.0213
12.001 | -34.5312 | -20.0302
12.0071 | -34.5316 | -20.0389
12.0132 | -34.5321 | -20.0477
12.0193 | -34.5325 | -20.0565
12.0254 | -34.5329 | -20.0654
12.0315 | -34.5334 | -20.0738
12.0376 | -34.5339 | -20.0825
12.0437 | -34.5343 | -20.0914
12.0498 | -34.5347 | -20.1003
12.0559 | -34.5352 | -20.1089
12.062 | -34.5357 | -20.1174
12.0681 | -34.5361 | -20.1263
12.0742 | -34.5366 | -20.135
12.0803 | -34.537 | -20.1439
12.0864 | -34.5375 | -20.1529
12.0925 | -34.538 | -20.1617
12.0986 | -34.5385 | -20.1704
12.1047 | -34.5389 | -20.1793
12.1109 | -34.5394 | -20.188
12.117 | -34.5398 | -20.1967
12.1231 | -34.5403 | -20.2054
12.1292 | -34.5408 | -20.2141
12.1353 | -34.5412 | -20.2228
12.1414 | -34.5417 | -20.2316
12.1475 | -34.5421 | -20.2406
12.1536 | -34.5426 | -20.2494
12.1597 | -34.543 | -20.2581
12.1658 | -34.5435 | -20.2669
12.1719 | -34.544 | -20.2754
12.178 | -34.5444 | -20.2841
12.1841 | -34.5449 | -20.293
12.1902 | -34.5453 | -20.3018
12.1963 | -34.5458 | -20.3105
12.2024 | -34.5463 | -20.319
12.2085 | -34.5468 | -20.3277
12.2146 | -34.5473 | -20.3365
12.2207 | -34.5477 | -20.3452
12.2268 | -34.5482 | -20.3538
12.2329 | -34.5486 | -20.3625
12.239 | -34.5491 | -20.3708
12.2451 | -34.5496 | -20.3795
12.2513 | -34.55 | -20.3883
12.2574 | -34.5505 | -20.397
12.2635 | -34.5509 | -20.4058
12.2696 | -34.5514 | -20.4144
12.2757 | -34.5518 | -20.4228
12.2818 | -34.5523 | -20.4316
12.2879 | -34.5528 | -20.4404
12.294 | -34.5533 | -20.449
12.3001 | -34.5537 | -20.4577
12.3062 | -34.5542 | -20.4664
12.3123 | -34.5547 | -20.4747
12.3184 | -34.5552 | -20.4833
12.3245 | -34.5556 | -20.4918
12.3306 | -34.5561 | -20.5005
12.3367 | -34.5565 | -20.5091
12.3428 | -34.557 | -20.5177
12.3489 | -34.5575 | -20.526
12.355 | -34.5579 | -20.5349
12.3611 | -34.5584 | -20.5436
12.3672 | -34.5588 | -20.5521
12.3733 | -34.5593 | -20.5607
12.3794 | -34.5598 | -20.569
12.3855 | -34.5602 | -20.5775
12.3916 | -34.5607 | -20.5861
12.3978 | -34.5612 | -20.5948
12.4039 | -34.5617 | -20.6035
12.41 | -34.5621 | -20.6122
12.4161 | -34.5626 | -20.6204
12.4222 | -34.5631 | -20.6289
12.4283 | -34.5635 | -20.6374
12.4344 | -34.564 | -20.6461
12.4405 | -34.5645 | -20.6548
12.4466 | -34.5649 | -20.6631
12.4527 | -34.5654 | -20.6717
12.4588 | -34.5658 | -20.6803
12.4649 | -34.5663 | -20.6889
12.471 | -34.5668 | -20.6976
12.4771 | -34.5672 | -20.7061
12.4832 | -34.5677 | -20.7148
12.4893 | -34.5681 | -20.7231
12.4954 | -34.5686 | -20.7316
12.5015 | -34.5691 | -20.7402
12.5076 | -34.5695 | -20.7488
12.5137 | -34.57 | -20.7574
12.5198 | -34.5705 | -20.7657
12.5259 | -34.5709 | -20.7742
12.532 | -34.5714 | -20.7828
12.5382 | -34.5719 | -20.7914
12.5443 | -34.5723 | -20.7999
12.5504 | -34.5728 | -20.8085
12.5565 | -34.5733 | -20.8169
12.5626 | -34.5738 | -20.8255
12.5687 | -34.5742 | -20.8341
12.5748 | -34.5747 | -20.8428
12.5809 | -34.5751 | -20.8513
12.587 | -34.5756 | -20.86
12.5931 | -34.5761 | -20.8683
12.5992 | -34.5766 | -20.8769
12.6053 | -34.5771 | -20.8854
12.6114 | -34.5776 | -20.8938
12.6175 | -34.578 | -20.9025
12.6236 | -34.5785 | -20.911
12.6297 | -34.5789 | -20.9194
12.6358 | -34.5794 | -20.9278
12.6419 | -34.5799 | -20.9362
12.648 | -34.5804 | -20.9448
12.6541 | -34.5808 | -20.9532
12.6602 | -34.5813 | -20.9615
12.6663 | -34.5817 | -20.9699
12.6724 | -34.5822 | -20.9783
12.6785 | -34.5827 | -20.9868
12.6847 | -34.5832 | -20.9953
12.6908 | -34.5836 | -21.0037
12.6969 | -34.5841 | -21.0119
12.703 | -34.5846 | -21.0205
12.7091 | -34.585 | -21.0291
12.7152 | -34.5855 | -21.0375
12.7213 | -34.5859 | -21.0461
12.7274 | -34.5864 | -21.0547
12.7335 | -34.5869 | -21.0629
12.7396 | -34.5873 | -21.0712
12.7457 | -34.5878 | -21.0798
12.7518 | -34.5882 | -21.0882
12.7579 | -34.5887 | -21.0969
12.764 | -34.5891 | -21.1054
12.7701 | -34.5896 | -21.1138
12.7762 | -34.5901 | -21.1222
12.7823 | -34.5906 | -21.1307
12.7884 | -34.5911 | -21.1392
12.7945 | -34.5915 | -21.1477
12.8006 | -34.592 | -21.156
12.8067 | -34.5924 | -21.1643
12.8128 | -34.5929 | -21.1727
12.8189 | -34.5934 | -21.1811
12.8251 | -34.5938 | -21.1896
12.8312 | -34.5943 | -21.1982
12.8373 | -34.5947 | -21.2065
12.8434 | -34.5952 | -21.2152
12.8495 | -34.5957 | -21.2236
12.8556 | -34.5961 | -21.232
12.8617 | -34.5966 | -21.2405
12.8678 | -34.5971 | -21.2492
12.8739 | -34.5975 | -21.2574
12.88 | -34.598 | -21.2658
12.8861 | -34.5984 | -21.2741
12.8922 | -34.5989 | -21.2826
12.8983 | -34.5993 | -21.291
12.9044 | -34.5998 | -21.2993
12.9105 | -34.6003 | -21.3077
12.9166 | -34.6007 | -21.3161
12.9227 | -34.6012 | -21.3244
12.9288 | -34.6016 | -21.3328
12.9349 | -34.6021 | -21.3412
12.941 | -34.6026 | -21.3496
12.9471 | -34.603 | -21.358
12.9532 | -34.6035 | -21.3666
12.9593 | -34.604 | -21.3751
12.9654 | -34.6045 | -21.3836
12.9716 | -34.6049 | -21.3922
12.9777 | -34.6054 | -21.4005
12.9838 | -34.6059 | -21.4087
12.9899 | -34.6063 | -21.4172
12.996 | -34.6068 | -21.4255
13.0021 | -34.6073 | -21.4338
13.0082 | -34.6077 | -21.4422
13.0143 | -34.6082 | -21.4503
13.0204 | -34.6086 | -21.4586
13.0265 | -34.6091 | -21.467
13.0326 | -34.6095 | -21.4755
13.0387 | -34.61 | -21.4839
13.0448 | -34.6105 | -21.4924
13.0509 | -34.6109 | -21.5008
13.057 | -34.6114 | -21.5091
13.0631 | -34.6119 | -21.5174
13.0692 | -34.6124 | -21.5258
13.0753 | -34.6128 | -21.5343
13.0814 | -34.6133 | -21.5428
13.0875 | -34.6138 | -21.5511
13.0936 | -34.6142 | -21.5596
13.0997 | -34.6147 | -21.5679
13.1058 | -34.6152 | -21.5762
13.112 | -34.6156 | -21.5846
13.1181 | -34.6161 | -21.5929
13.1242 | -34.6165 | -21.6014
13.1303 | -34.617 | -21.6098
13.1364 | -34.6175 | -21.6182
13.1425 | -34.6179 | -21.6265
13.1486 | -34.6184 | -21.6349
13.1547 | -34.6189 | -21.6434
13.1608 | -34.6193 | -21.6517
13.1669 | -34.6198 | -21.6599
13.173 | -34.6203 | -21.6684
13.1791 | -34.6207 | -21.6766
13.1852 | -34.6212 | -21.6849
13.1913 | -34.6217 | -21.6931
13.1974 | -34.6221 | -21.7016
13.2035 | -34.6225 | -21.71
13.2096 | -34.623 | -21.7183
13.2157 | -34.6235 | -21.7266
13.2218 | -34.624 | -21.735
13.2279 | -34.6244 | -21.7434
13.234 | -34.6249 | -21.7517
13.2401 | -34.6254 | -21.76
13.2462 | -34.6258 | -21.7683
13.2524 | -34.6263 | -21.7766
13.2585 | -34.6268 | -21.7849
13.2646 | -34.6272 | -21.7932
13.2707 | -34.6277 | -21.8016
13.2768 | -34.6282 | -21.8102
13.2829 | -34.6287 | -21.8186
13.289 | -34.6291 | -21.8268
13.2951 | -34.6296 | -21.8349
13.3012 | -34.6301 | -21.8432
13.3073 | -34.6306 | -21.8516
13.3134 | -34.631 | -21.8601
13.3195 | -34.6314 | -21.8683
13.3256 | -34.6319 | -21.8765
13.3317 | -34.6324 | -21.8849
13.3378 | -34.6328 | -21.8933
13.3439 | -34.6333 | -21.9018
13.35 | -34.6338 | -21.9103
13.3561 | -34.6342 | -21.9185
13.3622 | -34.6347 | -21.9268
13.3683 | -34.6352 | -21.9351
13.3744 | -34.6357 | -21.9434
13.3805 | -34.6361 | -21.9519
13.3866 | -34.6366 | -21.9601
13.3927 | -34.637 | -21.9684
13.3989 | -34.6375 | -21.9768
13.405 | -34.6379 | -21.9852
13.4111 | -34.6384 | -21.9935
13.4172 | -34.6388 | -22.0019
13.4233 | -34.6393 | -22.0104
13.4294 | -34.6398 | -22.019
13.4355 | -34.6402 | -22.0273
13.4416 | -34.6407 | -22.0356
13.4477 | -34.6412 | -22.0439
13.4538 | -34.6416 | -22.0522
13.4599 | -34.6421 | -22.0607
13.466 | -34.6426 | -22.0692
13.4721 | -34.643 | -22.0775
13.4782 | -34.6435 | -22.0858
13.4843 | -34.644 | -22.0943
13.4904 | -34.6444 | -22.1027
13.4965 | -34.6449 | -22.1112
13.5026 | -34.6453 | -22.1197
13.5087 | -34.6458 | -22.128
13.5148 | -34.6463 | -22.1366
13.5209 | -34.6467 | -22.1451
13.527 | -34.6472 | -22.1535
13.5331 | -34.6477 | -22.1617
13.5393 | -34.6481 | -22.1702
13.5454 | -34.6486 | -22.1786
13.5515 | -34.6491 | -22.1871
13.5576 | -34.6495 | -22.1955
13.5637 | -34.65 | -22.2039
13.5698 | -34.6505 | -22.2123
13.5759 | -34.6509 | -22.2204
13.582 | -34.6514 | -22.2287
13.5881 | -34.6519 | -22.2371
13.5942 | -34.6523 | -22.2456
13.6003 | -34.6528 | -22.2539
13.6064 | -34.6533 | -22.2625
13.6125 | -34.6538 | -22.271
13.6186 | -34.6542 | -22.2793
13.6247 | -34.6547 | -22.2876
13.6308 | -34.6552 | -22.296
13.6369 | -34.6556 | -22.3046
13.643 | -34.6561 | -22.3131
13.6491 | -34.6566 | -22.3217
13.6552 | -34.6571 | -22.3301
13.6613 | -34.6575 | -22.3384
13.6674 | -34.658 | -22.3468
13.6735 | -34.6584 | -22.3552
13.6796 | -34.6589 | -22.3636
13.6858 | -34.6594 | -22.372
13.6919 | -34.6599 | -22.3804
13.698 | -34.6603 | -22.3889
13.7041 | -34.6608 | -22.3973
13.7102 | -34.6613 | -22.4057
13.7163 | -34.6618 | -22.4141
13.7224 | -34.6622 | -22.4225
13.7285 | -34.6627 | -22.4309
13.7346 | -34.6632 | -22.4394
13.7407 | -34.6637 | -22.4477
13.7468 | -34.6642 | -22.4561
13.7529 | -34.6646 | -22.4644
13.759 | -34.6651 | -22.473
13.7651 | -34.6655 | -22.4814
13.7712 | -34.666 | -22.4899
13.7773 | -34.6665 | -22.4982
13.7834 | -34.6669 | -22.5066
13.7895 | -34.6674 | -22.5149
13.7956 | -34.6679 | -22.5234
13.8017 | -34.6684 | -22.5317
13.8078 | -34.6688 | -22.5402
13.8139 | -34.6693 | -22.5485
13.82 | -34.6698 | -22.5568
13.8262 | -34.6703 | -22.5653
13.8323 | -34.6708 | -22.5738
13.8384 | -34.6712 | -22.5823
13.8445 | -34.6717 | -22.5907
13.8506 | -34.6722 | -22.5992
13.8567 | -34.6726 | -22.6078
13.8628 | -34.6731 | -22.6164
13.8689 | -34.6736 | -22.6248
13.875 | -34.6741 | -22.6331
13.8811 | -34.6745 | -22.6416
13.8872 | -34.675 | -22.6499
13.8933 | -34.6755 | -22.6583
13.8994 | -34.6759 | -22.6668
13.9055 | -34.6764 | -22.6753
13.9116 | -34.6769 | -22.6837
13.9177 | -34.6773 | -22.6921
13.9238 | -34.6778 | -22.7005
13.9299 | -34.6783 | -22.7089
13.936 | -34.6787 | -22.7173
13.9421 | -34.6792 | -22.7257
13.9482 | -34.6797 | -22.734
13.9543 | -34.6801 | -22.7425
13.9604 | -34.6806 | -22.7509
13.9665 | -34.6811 | -22.7594
13.9727 | -34.6816 | -22.768
13.9788 | -34.6821 | -22.7763
13.9849 | -34.6826 | -22.7849
13.991 | -34.683 | -22.7933
13.9971 | -34.6835 | -22.8017
14.0032 | -34.684 | -22.8102
14.0093 | -34.6844 | -22.8186
14.0154 | -34.6849 | -22.827
14.0215 | -34.6854 | -22.8354
14.0276 | -34.6859 | -22.8437
14.0337 | -34.6864 | -22.8523
14.0398 | -34.6869 | -22.8607
14.0459 | -34.6873 | -22.8691
14.052 | -34.6878 | -22.8776
14.0581 | -34.6883 | -22.8861
14.0642 | -34.6887 | -22.8945
14.0703 | -34.6892 | -22.9029
14.0764 | -34.6897 | -22.9112
14.0825 | -34.6902 | -22.9197
14.0886 | -34.6907 | -22.9282
14.0947 | -34.6911 | -22.9366
14.1008 | -34.6916 | -22.945
14.1069 | -34.6921 | -22.9533
14.1131 | -34.6926 | -22.9619
14.1192 | -34.6931 | -22.9703
14.1253 | -34.6935 | -22.9786
14.1314 | -34.694 | -22.987
14.1375 | -34.6945 | -22.9954
14.1436 | -34.695 | -23.0039
14.1497 | -34.6955 | -23.0123
14.1558 | -34.696 | -23.0208
14.1619 | -34.6964 | -23.0291
14.168 | -34.6969 | -23.0374
14.1741 | -34.6974 | -23.0458
14.1802 | -34.6979 | -23.0542
14.1863 | -34.6984 | -23.0626
14.1924 | -34.6989 | -23.071
14.1985 | -34.6994 | -23.0795
14.2046 | -34.6998 | -23.088
14.2107 | -34.7003 | -23.0963
14.2168 | -34.7008 | -23.1048
14.2229 | -34.7013 | -23.1131
14.229 | -34.7018 | -23.1215
14.2351 | -34.7022 | -23.1298
14.2412 | -34.7027 | -23.1383
14.2473 | -34.7032 | -23.1467
14.2534 | -34.7036 | -23.155
14.2596 | -34.7041 | -23.1634
14.2657 | -34.7046 | -23.1717
14.2718 | -34.705 | -23.1802
14.2779 | -34.7055 | -23.1887
14.284 | -34.706 | -23.1971
14.2901 | -34.7065 | -23.2055
14.2962 | -34.707 | -23.214
14.3023 | -34.7074 | -23.2224
14.3084 | -34.7079 | -23.2308
14.3145 | -34.7084 | -23.2391
14.3206 | -34.7089 | -23.2477
14.3267 | -34.7094 | -23.256
14.3328 | -34.7098 | -23.2645
14.3389 | -34.7103 | -23.2728
14.345 | -34.7107 | -23.2812
14.3511 | -34.7112 | -23.2898
14.3572 | -34.7117 | -23.2984
14.3633 | -34.7122 | -23.3069
14.3694 | -34.7127 | -23.3154
14.3755 | -34.7131 | -23.3238
14.3816 | -34.7136 | -23.3325
14.3877 | -34.7141 | -23.3407
14.3938 | -34.7146 | -23.3494
14.4 | -34.7151 | -23.3579
14.4061 | -34.7156 | -23.3663
14.4122 | -34.716 | -23.3748
14.4183 | -34.7165 | -23.3831
14.4244 | -34.717 | -23.3916
14.4305 | -34.7175 | -23.4
14.4366 | -34.718 | -23.4084
14.4427 | -34.7184 | -23.4169
14.4488 | -34.7189 | -23.4254
14.4549 | -34.7194 | -23.4338
14.461 | -34.7199 | -23.4423
14.4671 | -34.7204 | -23.4509
14.4732 | -34.7209 | -23.4595
14.4793 | -34.7213 | -23.4681
14.4854 | -34.7218 | -23.4764
14.4915 | -34.7223 | -23.4846
14.4976 | -34.7228 | -23.4929
14.5037 | -34.7233 | -23.5013
14.5098 | -34.7238 | -23.5099
14.5159 | -34.7243 | -23.5183
14.522 | -34.7247 | -23.5265
14.5281 | -34.7252 | -23.5348
14.5342 | -34.7257 | -23.5433
14.5403 | -34.7262 | -23.5516
14.5465 | -34.7266 | -23.56
14.5526 | -34.7271 | -23.5686
14.5587 | -34.7276 | -23.5771
14.5648 | -34.7281 | -23.5853
14.5709 | -34.7286 | -23.5936
14.577 | -34.729 | -23.602
14.5831 | -34.7295 | -23.6105
14.5892 | -34.73 | -23.619
14.5953 | -34.7305 | -23.6272
14.6014 | -34.731 | -23.6356
14.6075 | -34.7314 | -23.6439
14.6136 | -34.7319 | -23.6523
14.6197 | -34.7323 | -23.6607
14.6258 | -34.7328 | -23.6693
14.6319 | -34.7333 | -23.6776
14.638 | -34.7338 | -23.6859
14.6441 | -34.7342 | -23.6944
14.6502 | -34.7347 | -23.7031
14.6563 | -34.7353 | -23.7116
14.6624 | -34.7357 | -23.72
14.6685 | -34.7362 | -23.7282
14.6746 | -34.7367 | -23.7366
14.6807 | -34.7372 | -23.7449
14.6869 | -34.7377 | -23.7533
14.693 | -34.7382 | -23.7616
14.6991 | -34.7386 | -23.7699
14.7052 | -34.7392 | -23.778
14.7113 | -34.7396 | -23.7865
14.7174 | -34.7401 | -23.7951
14.7235 | -34.7406 | -23.8034
14.7296 | -34.7411 | -23.8117
14.7357 | -34.7416 | -23.82
14.7418 | -34.7421 | -23.8286
14.7479 | -34.7425 | -23.8368
14.754 | -34.743 | -23.8453
14.7601 | -34.7435 | -23.8534
14.7662 | -34.744 | -23.8619
14.7723 | -34.7445 | -23.87
14.7784 | -34.7449 | -23.8784
14.7845 | -34.7454 | -23.8868
14.7906 | -34.7459 | -23.8952
14.7967 | -34.7463 | -23.9036
14.8028 | -34.7468 | -23.9121
14.8089 | -34.7473 | -23.9204
14.815 | -34.7478 | -23.9287
14.8211 | -34.7483 | -23.9372
14.8272 | -34.7487 | -23.9456
14.8334 | -34.7493 | -23.9539
14.8395 | -34.7498 | -23.9621
14.8456 | -34.7502 | -23.9704
14.8517 | -34.7507 | -23.9789
14.8578 | -34.7512 | -23.9873
14.8639 | -34.7517 | -23.9955
14.87 | -34.7522 | -24.0039
14.8761 | -34.7526 | -24.0122
14.8822 | -34.7531 | -24.0206
14.8883 | -34.7536 | -24.0291
14.8944 | -34.7541 | -24.0374
14.9005 | -34.7545 | -24.0458
14.9066 | -34.755 | -24.0541
14.9127 | -34.7555 | -24.0623
14.9188 | -34.756 | -24.0707
14.9249 | -34.7564 | -24.079
14.931 | -34.7569 | -24.0874
14.9371 | -34.7574 | -24.0957
14.9432 | -34.7579 | -24.1039
14.9493 | -34.7584 | -24.1122
14.9554 | -34.7589 | -24.1206
14.9615 | -34.7593 | -24.1288
14.9676 | -34.7598 | -24.1372
14.9738 | -34.7603 | -24.1454
14.9799 | -34.7608 | -24.1538
14.986 | -34.7613 | -24.162
14.9921 | -34.7617 | -24.1702
14.9982 | -34.7622 | -24.1784
15.0043 | -34.7627 | -24.1867
15.0104 | -34.7632 | -24.1951
15.0165 | -34.7637 | -24.2038
15.0226 | -34.7642 | -24.212
15.0287 | -34.7646 | -24.2203
15.0348 | -34.7651 | -24.2287
15.0409 | -34.7656 | -24.2369
15.047 | -34.7661 | -24.2452
15.0531 | -34.7666 | -24.2533
15.0592 | -34.7671 | -24.2617
15.0653 | -34.7676 | -24.27
15.0714 | -34.7681 | -24.2784
15.0775 | -34.7685 | -24.2866
15.0836 | -34.769 | -24.2948
15.0897 | -34.7695 | -24.3029
15.0958 | -34.77 | -24.3112
15.1019 | -34.7705 | -24.3196
15.108 | -34.7709 | -24.3278
15.1141 | -34.7714 | -24.3363
15.1203 | -34.7719 | -24.3447
15.1264 | -34.7724 | -24.3528
15.1325 | -34.7729 | -24.3611
15.1386 | -34.7733 | -24.3693
15.1447 | -34.7738 | -24.3776
15.1508 | -34.7743 | -24.3858
15.1569 | -34.7748 | -24.3941
15.163 | -34.7753 | -24.4021
15.1691 | -34.7758 | -24.4102
15.1752 | -34.7763 | -24.4185
15.1813 | -34.7767 | -24.4267
15.1874 | -34.7772 | -24.435
15.1935 | -34.7777 | -24.443
15.1996 | -34.7782 | -24.4512
15.2057 | -34.7787 | -24.4594
15.2118 | -34.7791 | -24.4678
15.2179 | -34.7796 | -24.4761
15.224 | -34.7801 | -24.4843
15.2301 | -34.7806 | -24.4923
15.2362 | -34.7811 | -24.5006
15.2423 | -34.7816 | -24.5089
15.2484 | -34.7821 | -24.5172
15.2545 | -34.7825 | -24.5254
15.2607 | -34.783 | -24.5336
15.2668 | -34.7835 | -24.5416
15.2729 | -34.784 | -24.55
15.279 | -34.7845 | -24.5584
15.2851 | -34.785 | -24.5666
15.2912 | -34.7855 | -24.5749
15.2973 | -34.786 | -24.5832
15.3034 | -34.7864 | -24.5913
15.3095 | -34.7869 | -24.5993
15.3156 | -34.7874 | -24.6075
15.3217 | -34.7879 | -24.6155
15.3278 | -34.7884 | -24.6239
15.3339 | -34.7889 | -24.6319
15.34 | -34.7894 | -24.6397
15.3461 | -34.7899 | -24.648
15.3522 | -34.7903 | -24.656
15.3583 | -34.7908 | -24.6644
15.3644 | -34.7913 | -24.6723
15.3705 | -34.7918 | -24.6802
15.3766 | -34.7922 | -24.6884
15.3827 | -34.7927 | -24.6964
15.3888 | -34.7932 | -24.7045
15.3949 | -34.7937 | -24.7127
15.401 | -34.7942 | -24.7211
15.4072 | -34.7947 | -24.7289
15.4133 | -34.7951 | -24.7373
15.4194 | -34.7956 | -24.7454
15.4255 | -34.7961 | -24.7535
15.4316 | -34.7965 | -24.7617
15.4377 | -34.797 | -24.7699
15.4438 | -34.7975 | -24.7779
15.4499 | -34.798 | -24.7862
15.456 | -34.7985 | -24.7941
15.4621 | -34.7989 | -24.8024
15.4682 | -34.7994 | -24.8108
15.4743 | -34.7999 | -24.8185
15.4804 | -34.8004 | -24.8267
15.4865 | -34.8008 | -24.8349
15.4926 | -34.8013 | -24.8431
15.4987 | -34.8018 | -24.8512
15.5048 | -34.8023 | -24.8593
15.5109 | -34.8028 | -24.8672
15.517 | -34.8033 | -24.8755
15.5231 | -34.8038 | -24.8836
15.5292 | -34.8043 | -24.8916
15.5353 | -34.8048 | -24.9001
15.5414 | -34.8053 | -24.9083
15.5476 | -34.8058 | -24.9161
15.5537 | -34.8063 | -24.9245
15.5598 | -34.8068 | -24.9328
15.5659 | -34.8073 | -24.9409
15.572 | -34.8078 | -24.9491
15.5781 | -34.8083 | -24.9573
15.5842 | -34.8088 | -24.9652
15.5903 | -34.8092 | -24.9732
15.5964 | -34.8097 | -24.9811
15.6025 | -34.8102 | -24.989
15.6086 | -34.8107 | -24.9974
15.6147 | -34.8111 | -25.0056
15.6208 | -34.8116 | -25.0133
15.6269 | -34.8121 | -25.0213
15.633 | -34.8126 | -25.0295
15.6391 | -34.8131 | -25.0376
15.6452 | -34.8136 | -25.0456
15.6513 | -34.8141 | -25.0535
15.6574 | -34.8145 | -25.0617
15.6635 | -34.815 | -25.0697
15.6696 | -34.8155 | -25.0777
15.6757 | -34.816 | -25.0858
15.6818 | -34.8165 | -25.094
15.688 | -34.817 | -25.1019
15.6941 | -34.8174 | -25.1099
15.7002 | -34.8179 | -25.1179
15.7063 | -34.8184 | -25.1258
15.7124 | -34.8189 | -25.1339
15.7185 | -34.8193 | -25.142
15.7246 | -34.8198 | -25.1498
15.7307 | -34.8203 | -25.1576
15.7368 | -34.8208 | -25.1657
15.7429 | -34.8213 | -25.1736
15.749 | -34.8218 | -25.1817
15.7551 | -34.8223 | -25.1898
15.7612 | -34.8228 | -25.1976
15.7673 | -34.8233 | -25.2056
15.7734 | -34.8238 | -25.2137
15.7795 | -34.8243 | -25.2219
15.7856 | -34.8247 | -25.2298
15.7917 | -34.8252 | -25.2374
15.7978 | -34.8257 | -25.2455
15.8039 | -34.8262 | -25.2536
15.81 | -34.8266 | -25.2613
15.8161 | -34.8271 | -25.2694
15.8222 | -34.8276 | -25.2772
15.8283 | -34.8281 | -25.2848
15.8345 | -34.8286 | -25.2928
15.8406 | -34.829 | -25.3009
15.8467 | -34.8295 | -25.3091
15.8528 | -34.83 | -25.317
15.8589 | -34.8305 | -25.3251
15.865 | -34.831 | -25.3327
15.8711 | -34.8315 | -25.3407
15.8772 | -34.832 | -25.3484
15.8833 | -34.8324 | -25.3564
15.8894 | -34.8329 | -25.3646
15.8955 | -34.8334 | -25.3728
15.9016 | -34.8338 | -25.3805
15.9077 | -34.8343 | -25.3885
15.9138 | -34.8348 | -25.3966
15.9199 | -34.8353 | -25.4043
15.926 | -34.8358 | -25.4123
15.9321 | -34.8363 | -25.4202
15.9382 | -34.8367 | -25.4279
15.9443 | -34.8372 | -25.436
15.9504 | -34.8377 | -25.4439
15.9565 | -34.8382 | -25.4519
15.9626 | -34.8387 | -25.46
15.9687 | -34.8392 | -25.4676
15.9749 | -34.8396 | -25.4752
15.981 | -34.8401 | -25.4831
15.9871 | -34.8406 | -25.4912
15.9932 | -34.8411 | -25.4989
15.9993 | -34.8415 | -25.5067
16.0054 | -34.842 | -25.5141
16.0115 | -34.8425 | -25.5222
16.0176 | -34.843 | -25.5305
16.0237 | -34.8435 | -25.5385
16.0298 | -34.844 | -25.5465
16.0359 | -34.8444 | -25.5544
16.042 | -34.845 | -25.5621
16.0481 | -34.8455 | -25.5699
16.0542 | -34.8459 | -25.578
16.0603 | -34.8464 | -25.5856
16.0664 | -34.8468 | -25.5936
16.0725 | -34.8473 | -25.6013
16.0786 | -34.8478 | -25.6091
16.0847 | -34.8483 | -25.6171
16.0908 | -34.8488 | -25.6251
16.0969 | -34.8493 | -25.6332
16.103 | -34.8497 | -25.641
16.1091 | -34.8502 | -25.6487
16.1152 | -34.8507 | -25.6567
16.1214 | -34.8512 | -25.6647
16.1275 | -34.8517 | -25.6727
16.1336 | -34.8521 | -25.6807
16.1397 | -34.8526 | -25.6884
16.1458 | -34.8531 | -25.696
16.1519 | -34.8537 | -25.7039
16.158 | -34.8541 | -25.7119
16.1641 | -34.8545 | -25.7199
16.1702 | -34.855 | -25.7278
16.1763 | -34.8555 | -25.7359
16.1824 | -34.856 | -25.7437
16.1885 | -34.8564 | -25.7515
16.1946 | -34.8569 | -25.7592
16.2007 | -34.8574 | -25.767
16.2068 | -34.8579 | -25.7748
16.2129 | -34.8584 | -25.7824
16.219 | -34.8589 | -25.79
16.2251 | -34.8594 | -25.7978
16.2312 | -34.8599 | -25.8054
16.2373 | -34.8603 | -25.8131
16.2434 | -34.8608 | -25.8208
16.2495 | -34.8613 | -25.8285
16.2556 | -34.8618 | -25.8365
16.2618 | -34.8623 | -25.8445
16.2679 | -34.8628 | -25.8525
16.274 | -34.8633 | -25.8605
16.2801 | -34.8637 | -25.8682
16.2862 | -34.8642 | -25.8758
16.2923 | -34.8647 | -25.8839
16.2984 | -34.8651 | -25.8918
16.3045 | -34.8656 | -25.8998
16.3106 | -34.8661 | -25.9078
16.3167 | -34.8666 | -25.9157
16.3228 | -34.867 | -25.9232
16.3289 | -34.8675 | -25.9311
16.335 | -34.868 | -25.9387
16.3411 | -34.8685 | -25.9465
16.3472 | -34.869 | -25.9548
16.3533 | -34.8695 | -25.9628
16.3594 | -34.87 | -25.9703
16.3655 | -34.8705 | -25.978
16.3716 | -34.871 | -25.9856
16.3777 | -34.8714 | -25.9933
16.3838 | -34.8719 | -26.001
16.3899 | -34.8724 | -26.0086
16.396 | -34.8729 | -26.0165
16.4021 | -34.8733 | -26.024
16.4083 | -34.8738 | -26.0317
16.4144 | -34.8743 | -26.0393
16.4205 | -34.8748 | -26.0469
16.4266 | -34.8753 | -26.0543
16.4327 | -34.8758 | -26.0619
16.4388 | -34.8762 | -26.0698
16.4449 | -34.8767 | -26.0775
16.451 | -34.8772 | -26.0852
16.4571 | -34.8776 | -26.0932
16.4632 | -34.8781 | -26.1006
16.4693 | -34.8785 | -26.1085
16.4754 | -34.879 | -26.1164
16.4815 | -34.8795 | -26.1241
16.4876 | -34.88 | -26.132
16.4937 | -34.8805 | -26.1396
16.4998 | -34.8809 | -26.1471
16.5059 | -34.8814 | -26.1547
16.512 | -34.8819 | -26.1627
16.5181 | -34.8824 | -26.1705
16.5242 | -34.8829 | -26.1784
16.5303 | -34.8833 | -26.1859
16.5364 | -34.8838 | -26.1936
16.5425 | -34.8843 | -26.2014
16.5487 | -34.8848 | -26.2091
16.5548 | -34.8853 | -26.2169
16.5609 | -34.8858 | -26.2248
16.567 | -34.8863 | -26.2321
16.5731 | -34.8868 | -26.24
16.5792 | -34.8872 | -26.2477
16.5853 | -34.8877 | -26.2554
16.5914 | -34.8882 | -26.2631
16.5975 | -34.8887 | -26.271
16.6036 | -34.8891 | -26.2785
16.6097 | -34.8896 | -26.2864
16.6158 | -34.8901 | -26.2941
16.6219 | -34.8906 | -26.3018
16.628 | -34.891 | -26.3095
16.6341 | -34.8915 | -26.3172
16.6402 | -34.892 | -26.3245
16.6463 | -34.8924 | -26.3323
16.6524 | -34.8929 | -26.3399
16.6585 | -34.8934 | -26.3476
16.6646 | -34.8939 | -26.3552
16.6707 | -34.8943 | -26.3627
16.6768 | -34.8948 | -26.3704
16.6829 | -34.8953 | -26.378
16.689 | -34.8957 | -26.3857
16.6952 | -34.8962 | -26.3934
16.7013 | -34.8967 | -26.4009
16.7074 | -34.8972 | -26.4083
16.7135 | -34.8977 | -26.4161
16.7196 | -34.8981 | -26.4238
16.7257 | -34.8986 | -26.4314
16.7318 | -34.8991 | -26.4389
16.7379 | -34.8996 | -26.4465
16.744 | -34.9 | -26.454
16.7501 | -34.9005 | -26.4616
16.7562 | -34.9009 | -26.4692
16.7623 | -34.9014 | -26.4768
16.7684 | -34.9019 | -26.4844
16.7745 | -34.9023 | -26.4921
16.7806 | -34.9028 | -26.4993
16.7867 | -34.9033 | -26.5069
16.7928 | -34.9038 | -26.5146
16.7989 | -34.9042 | -26.5221
16.805 | -34.9047 | -26.5297
16.8111 | -34.9052 | -26.5371
16.8172 | -34.9057 | -26.5448
16.8233 | -34.9062 | -26.5524
16.8294 | -34.9066 | -26.56
16.8356 | -34.9071 | -26.5678
16.8417 | -34.9076 | -26.5752
16.8478 | -34.9081 | -26.5827
16.8539 | -34.9085 | -26.5901
16.86 | -34.909 | -26.5978
16.8661 | -34.9095 | -26.6053
16.8722 | -34.91 | -26.613
16.8783 | -34.9104 | -26.6208
16.8844 | -34.9109 | -26.6281
16.8905 | -34.9114 | -26.6358
16.8966 | -34.9119 | -26.6435
16.9027 | -34.9124 | -26.651
16.9088 | -34.9128 | -26.6587
16.9149 | -34.9133 | -26.6664
16.921 | -34.9138 | -26.6738
16.9271 | -34.9143 | -26.6815
16.9332 | -34.9147 | -26.6892
16.9393 | -34.9152 | -26.6968
16.9454 | -34.9157 | -26.7043
16.9515 | -34.9162 | -26.712
16.9576 | -34.9166 | -26.7193
16.9637 | -34.9171 | -26.7268
16.9698 | -34.9176 | -26.7344
16.9759 | -34.9181 | -26.7418
16.9821 | -34.9186 | -26.7496
16.9882 | -34.919 | -26.7569
16.9943 | -34.9195 | -26.7645
17.0004 | -34.92 | -26.7722
17.0065 | -34.9205 | -26.7799
17.0126 | -34.921 | -26.7873
17.0187 | -34.9214 | -26.7951
17.0248 | -34.9219 | -26.8025
17.0309 | -34.9223 | -26.8101
17.037 | -34.9228 | -26.8176
17.0431 | -34.9233 | -26.8252
17.0492 | -34.9237 | -26.8327
17.0553 | -34.9242 | -26.8403
17.0614 | -34.9247 | -26.8477
17.0675 | -34.9252 | -26.8552
17.0736 | -34.9256 | -26.8629
17.0797 | -34.9261 | -26.8704
17.0858 | -34.9266 | -26.8781
17.0919 | -34.9271 | -26.8857
17.098 | -34.9276 | -26.893
17.1041 | -34.928 | -26.9006
17.1102 | -34.9285 | -26.9081
17.1163 | -34.929 | -26.9156
17.1225 | -34.9295 | -26.923
17.1286 | -34.93 | -26.9303
17.1347 | -34.9305 | -26.9379
17.1408 | -34.9309 | -26.9455
17.1469 | -34.9314 | -26.9528
17.153 | -34.9319 | -26.9603
17.1591 | -34.9323 | -26.9678
17.1652 | -34.9328 | -26.9752
17.1713 | -34.9333 | -26.9827
17.1774 | -34.9338 | -26.9902
17.1835 | -34.9343 | -26.9975
17.1896 | -34.9348 | -27.005
17.1957 | -34.9353 | -27.0125
17.2018 | -34.9357 | -27.0197
17.2079 | -34.9362 | -27.0272
17.214 | -34.9367 | -27.0348
17.2201 | -34.9372 | -27.0424
17.2262 | -34.9376 | -27.05
17.2323 | -34.9381 | -27.0576
17.2384 | -34.9386 | -27.0649
17.2445 | -34.9391 | -27.0723
17.2506 | -34.9396 | -27.0798
17.2567 | -34.94 | -27.0874
17.2628 | -34.9405 | -27.0947
17.269 | -34.941 | -27.1021
17.2751 | -34.9415 | -27.1096
17.2812 | -34.9419 | -27.1172
17.2873 | -34.9424 | -27.1247
17.2934 | -34.9428 | -27.1323
17.2995 | -34.9433 | -27.1399
17.3056 | -34.9438 | -27.1471
17.3117 | -34.9443 | -27.1547
17.3178 | -34.9447 | -27.1621
17.3239 | -34.9452 | -27.1696
17.33 | -34.9456 | -27.1771
17.3361 | -34.9461 | -27.1845
17.3422 | -34.9466 | -27.1918
17.3483 | -34.9471 | -27.1993
17.3544 | -34.9475 | -27.2068
17.3605 | -34.948 | -27.2143
17.3666 | -34.9485 | -27.2218
17.3727 | -34.9489 | -27.2293
17.3788 | -34.9494 | -27.2363
17.3849 | -34.9499 | -27.2436
17.391 | -34.9504 | -27.2511
17.3971 | -34.9508 | -27.2585
17.4032 | -34.9513 | -27.2658
17.4094 | -34.9518 | -27.2732
17.4155 | -34.9522 | -27.2807
17.4216 | -34.9527 | -27.2881
17.4277 | -34.9532 | -27.2956
17.4338 | -34.9536 | -27.3031
17.4399 | -34.9541 | -27.3106
17.446 | -34.9546 | -27.3178
17.4521 | -34.9551 | -27.325
17.4582 | -34.9556 | -27.3324
17.4643 | -34.956 | -27.3397
17.4704 | -34.9565 | -27.3471
17.4765 | -34.957 | -27.3546
17.4826 | -34.9575 | -27.3618
17.4887 | -34.958 | -27.3693
17.4948 | -34.9584 | -27.3767
17.5009 | -34.9589 | -27.3841
17.507 | -34.9594 | -27.3917
17.5131 | -34.9599 | -27.3992
17.5192 | -34.9603 | -27.4063
17.5253 | -34.9608 | -27.4138
17.5314 | -34.9613 | -27.4212
17.5375 | -34.9617 | -27.4287
17.5436 | -34.9622 | -27.4362
17.5497 | -34.9627 | -27.4437
17.5559 | -34.9632 | -27.4509
17.562 | -34.9636 | -27.4582
17.5681 | -34.9641 | -27.4656
17.5742 | -34.9646 | -27.4732
17.5803 | -34.9651 | -27.4806
17.5864 | -34.9656 | -27.4879
17.5925 | -34.966 | -27.4954
17.5986 | -34.9665 | -27.5027
17.6047 | -34.967 | -27.5103
17.6108 | -34.9675 | -27.5178
17.6169 | -34.9679 | -27.5252
17.623 | -34.9684 | -27.5324
17.6291 | -34.9689 | -27.5396
17.6352 | -34.9694 | -27.5471
17.6413 | -34.9698 | -27.5545
17.6474 | -34.9703 | -27.5621
17.6535 | -34.9707 | -27.5696
17.6596 | -34.9712 | -27.5767
17.6657 | -34.9717 | -27.5841
17.6718 | -34.9722 | -27.5915
17.6779 | -34.9726 | -27.5989
17.684 | -34.9731 | -27.6063
17.6901 | -34.9736 | -27.6136
17.6963 | -34.974 | -27.6208
17.7024 | -34.9745 | -27.628
17.7085 | -34.975 | -27.6353
17.7146 | -34.9754 | -27.6426
17.7207 | -34.9759 | -27.6499
17.7268 | -34.9764 | -27.6571
17.7329 | -34.9769 | -27.6645
17.739 | -34.9773 | -27.6718
17.7451 | -34.9778 | -27.6791
17.7512 | -34.9783 | -27.6865
17.7573 | -34.9788 | -27.6939
17.7634 | -34.9792 | -27.7009
17.7695 | -34.9797 | -27.7082
17.7756 | -34.9802 | -27.7155
17.7817 | -34.9806 | -27.7229
17.7878 | -34.9811 | -27.7304
17.7939 | -34.9816 | -27.7378
17.8 | -34.9821 | -27.7449
17.8061 | -34.9825 | -27.7523
17.8122 | -34.983 | -27.7597
17.8183 | -34.9835 | -27.767
17.8244 | -34.9839 | -27.7744
17.8305 | -34.9844 | -27.7817
17.8366 | -34.9849 | -27.7888
17.8428 | -34.9853 | -27.7963
17.8489 | -34.9858 | -27.8035
17.855 | -34.9863 | -27.8108
17.8611 | -34.9867 | -27.8181
17.8672 | -34.9872 | -27.8254
17.8733 | -34.9877 | -27.8327
17.8794 | -34.9882 | -27.84
17.8855 | -34.9886 | -27.8472
17.8916 | -34.9891 | -27.8547
17.8977 | -34.9895 | -27.862
17.9038 | -34.99 | -27.8691
17.9099 | -34.9905 | -27.8765
17.916 | -34.991 | -27.8839
17.9221 | -34.9914 | -27.8913
17.9282 | -34.9919 | -27.8986
17.9343 | -34.9924 | -27.9059
17.9404 | -34.9928 | -27.9131
17.9465 | -34.9933 | -27.9203
17.9526 | -34.9938 | -27.9278
17.9587 | -34.9943 | -27.9351
17.9648 | -34.9947 | -27.9423
17.9709 | -34.9952 | -27.9496
17.977 | -34.9957 | -27.9569
17.9832 | -34.9961 | -27.9642
17.9893 | -34.9966 | -27.9716
17.9954 | -34.9971 | -27.9789
18.0015 | -34.9975 | -27.9862
18.0076 | -34.998 | -27.9933
18.0137 | -34.9984 | -28.0006
18.0198 | -34.9989 | -28.0079
18.0259 | -34.9993 | -28.0152
18.032 | -34.9998 | -28.0226
18.0381 | -35.0003 | -28.03
18.0442 | -35.0008 | -28.0371
18.0503 | -35.0012 | -28.0445
18.0564 | -35.0017 | -28.0517
18.0625 | -35.0022 | -28.0591
18.0686 | -35.0026 | -28.0665
18.0747 | -35.0031 | -28.0739
18.0808 | -35.0036 | -28.0809
18.0869 | -35.004 | -28.0882
18.093 | -35.0045 | -28.0956
18.0991 | -35.005 | -28.1029
18.1052 | -35.0054 | -28.1101
18.1113 | -35.0059 | -28.1174
18.1174 | -35.0064 | -28.1247
18.1236 | -35.0069 | -28.1319
18.1297 | -35.0074 | -28.1392
18.1358 | -35.0078 | -28.1467
18.1419 | -35.0083 | -28.1539
18.148 | -35.0088 | -28.1611
18.1541 | -35.0093 | -28.1681
18.1602 | -35.0098 | -28.1755
18.1663 | -35.0102 | -28.1827
18.1724 | -35.0107 | -28.19
18.1785 | -35.0112 | -28.1971
18.1846 | -35.0116 | -28.2045
18.1907 | -35.0121 | -28.2117
18.1968 | -35.0125 | -28.2191
18.2029 | -35.013 | -28.2264
18.209 | -35.0135 | -28.2338
18.2151 | -35.014 | -28.2411
18.2212 | -35.0144 | -28.2483
18.2273 | -35.0149 | -28.2555
18.2334 | -35.0154 | -28.2626
18.2395 | -35.0158 | -28.2698
18.2456 | -35.0163 | -28.2772
18.2517 | -35.0167 | -28.2844
18.2578 | -35.0172 | -28.2914
18.2639 | -35.0177 | -28.2986
18.2701 | -35.0182 | -28.3059
18.2762 | -35.0186 | -28.313
18.2823 | -35.0191 | -28.32
18.2884 | -35.0196 | -28.3273
18.2945 | -35.0201 | -28.3343
18.3006 | -35.0205 | -28.3417
18.3067 | -35.021 | -28.3488
18.3128 | -35.0215 | -28.3563
18.3189 | -35.0219 | -28.3636
18.325 | -35.0224 | -28.3707
18.3311 | -35.0228 | -28.378
18.3372 | -35.0233 | -28.3853
18.3433 | -35.0238 | -28.3925
18.3494 | -35.0243 | -28.3997
18.3555 | -35.0248 | -28.407
18.3616 | -35.0252 | -28.414
18.3677 | -35.0257 | -28.4213
18.3738 | -35.0262 | -28.4284
18.3799 | -35.0267 | -28.4357
18.386 | -35.0271 | -28.4429
18.3921 | -35.0275 | -28.4502
18.3982 | -35.028 | -28.4572
18.4043 | -35.0284 | -28.4645
18.4105 | -35.0289 | -28.4716
18.4166 | -35.0293 | -28.4789
18.4227 | -35.0298 | -28.4861
18.4288 | -35.0303 | -28.4934
18.4349 | -35.0308 | -28.5004
18.441 | -35.0312 | -28.5076
18.4471 | -35.0317 | -28.5148
18.4532 | -35.0322 | -28.5222
18.4593 | -35.0326 | -28.5291
18.4654 | -35.0331 | -28.5364
18.4715 | -35.0336 | -28.5439
18.4776 | -35.034 | -28.5513
18.4837 | -35.0345 | -28.5586
18.4898 | -35.0349 | -28.5659
18.4959 | -35.0354 | -28.5732
18.502 | -35.0359 | -28.5804
18.5081 | -35.0363 | -28.5877
18.5142 | -35.0368 | -28.5948
18.5203 | -35.0372 | -28.6019
18.5264 | -35.0377 | -28.6091
18.5325 | -35.0382 | -28.6162
18.5386 | -35.0386 | -28.6234
18.5447 | -35.0391 | -28.6306
18.5508 | -35.0396 | -28.6376
18.557 | -35.04 | -28.645
18.5631 | -35.0405 | -28.6522
18.5692 | -35.041 | -28.6593
18.5753 | -35.0414 | -28.6664
18.5814 | -35.0419 | -28.6735
18.5875 | -35.0424 | -28.6807
18.5936 | -35.0428 | -28.6877
18.5997 | -35.0433 | -28.695
18.6058 | -35.0437 | -28.7018
18.6119 | -35.0442 | -28.7089
18.618 | -35.0447 | -28.7162
18.6241 | -35.0451 | -28.7233
18.6302 | -35.0456 | -28.7305
18.6363 | -35.046 | -28.7377
18.6424 | -35.0465 | -28.7448
18.6485 | -35.047 | -28.7519
18.6546 | -35.0474 | -28.7591
18.6607 | -35.0479 | -28.7663
18.6668 | -35.0483 | -28.7735
18.6729 | -35.0488 | -28.7808
18.679 | -35.0493 | -28.7877
18.6851 | -35.0498 | -28.7948
18.6912 | -35.0502 | -28.8019
18.6974 | -35.0506 | -28.809
18.7035 | -35.0511 | -28.8162
18.7096 | -35.0516 | -28.8233
18.7157 | -35.052 | -28.8305
18.7218 | -35.0525 | -28.8375
18.7279 | -35.053 | -28.8448
18.734 | -35.0535 | -28.8521
18.7401 | -35.0539 | -28.8592
18.7462 | -35.0544 | -28.8661
18.7523 | -35.0549 | -28.8734
18.7584 | -35.0553 | -28.8805
18.7645 | -35.0558 | -28.8876
18.7706 | -35.0563 | -28.8947
18.7767 | -35.0567 | -28.9019
18.7828 | -35.0572 | -28.9088
18.7889 | -35.0576 | -28.9159
18.795 | -35.0581 | -28.9232
18.8011 | -35.0586 | -28.9301
18.8072 | -35.0591 | -28.9374
18.8133 | -35.0595 | -28.9445
18.8194 | -35.06 | -28.9517
18.8255 | -35.0605 | -28.9588
18.8316 | -35.0609 | -28.966
18.8377 | -35.0613 | -28.9732
18.8439 | -35.0618 | -28.9802
18.85 | -35.0622 | -28.9874
18.8561 | -35.0627 | -28.9943
18.8622 | -35.0632 | -29.0015
18.8683 | -35.0636 | -29.0085
18.8744 | -35.0641 | -29.0156
18.8805 | -35.0645 | -29.023
18.8866 | -35.065 | -29.0302
18.8927 | -35.0655 | -29.0369
18.8988 | -35.066 | -29.0441
18.9049 | -35.0664 | -29.0512
18.911 | -35.0669 | -29.0584
18.9171 | -35.0674 | -29.0655
18.9232 | -35.0678 | -29.0725
18.9293 | -35.0683 | -29.0796
18.9354 | -35.0688 | -29.0866
18.9415 | -35.0693 | -29.0938
18.9476 | -35.0697 | -29.1008
18.9537 | -35.0702 | -29.108
18.9598 | -35.0707 | -29.1149
18.9659 | -35.0711 | -29.1218
18.972 | -35.0716 | -29.1288
18.9781 | -35.072 | -29.136
18.9843 | -35.0725 | -29.143
18.9904 | -35.073 | -29.15
18.9965 | -35.0735 | -29.1568
19.0026 | -35.0739 | -29.1639
19.0087 | -35.0744 | -29.171
19.0148 | -35.0749 | -29.178
19.0209 | -35.0753 | -29.1853
19.027 | -35.0758 | -29.1925
19.0331 | -35.0763 | -29.1995
19.0392 | -35.0768 | -29.2067
19.0453 | -35.0772 | -29.2137
19.0514 | -35.0777 | -29.221
19.0575 | -35.0782 | -29.2281
19.0636 | -35.0786 | -29.235
19.0697 | -35.0791 | -29.2422
19.0758 | -35.0796 | -29.2493
19.0819 | -35.08 | -29.2562
19.088 | -35.0805 | -29.2632
19.0941 | -35.0809 | -29.2705
19.1002 | -35.0814 | -29.2775
19.1063 | -35.0819 | -29.2846
19.1124 | -35.0823 | -29.2917
19.1185 | -35.0828 | -29.2988
19.1246 | -35.0833 | -29.3059
19.1308 | -35.0837 | -29.3129
19.1369 | -35.0842 | -29.3198
19.143 | -35.0847 | -29.3269
19.1491 | -35.0851 | -29.3341
19.1552 | -35.0856 | -29.3414
19.1613 | -35.086 | -29.3485
19.1674 | -35.0865 | -29.3556
19.1735 | -35.0869 | -29.3625
19.1796 | -35.0874 | -29.3696
19.1857 | -35.0878 | -29.3768
19.1918 | -35.0883 | -29.3838
19.1979 | -35.0888 | -29.3911
19.204 | -35.0893 | -29.3981
19.2101 | -35.0897 | -29.405
19.2162 | -35.0902 | -29.4118
19.2223 | -35.0906 | -29.4189
19.2284 | -35.0911 | -29.426
19.2345 | -35.0916 | -29.4332
19.2406 | -35.092 | -29.4401
19.2467 | -35.0925 | -29.447
19.2528 | -35.0929 | -29.4541
19.2589 | -35.0934 | -29.4611
19.265 | -35.0939 | -29.4682
19.2712 | -35.0944 | -29.4754
19.2773 | -35.0948 | -29.4825
19.2834 | -35.0953 | -29.4894
19.2895 | -35.0957 | -29.4966
19.2956 | -35.0962 | -29.5036
19.3017 | -35.0967 | -29.5106
19.3078 | -35.0971 | -29.5177
19.3139 | -35.0976 | -29.5247
19.32 | -35.0981 | -29.5317
19.3261 | -35.0985 | -29.5387
19.3322 | -35.099 | -29.5457
19.3383 | -35.0995 | -29.5526
19.3444 | -35.0999 | -29.5594
19.3505 | -35.1004 | -29.5665
19.3566 | -35.1008 | -29.5735
19.3627 | -35.1013 | -29.5805
19.3688 | -35.1018 | -29.5875
19.3749 | -35.1023 | -29.5945
19.381 | -35.1028 | -29.6012
19.3871 | -35.1032 | -29.6082
19.3932 | -35.1037 | -29.6151
19.3993 | -35.1042 | -29.6221
19.4054 | -35.1046 | -29.6291
19.4115 | -35.1051 | -29.6363
19.4177 | -35.1056 | -29.6431
19.4238 | -35.106 | -29.65
19.4299 | -35.1065 | -29.657
19.436 | -35.1069 | -29.6641
19.4421 | -35.1074 | -29.6711
19.4482 | -35.1078 | -29.6781
19.4543 | -35.1083 | -29.6849
19.4604 | -35.1088 | -29.6918
19.4665 | -35.1092 | -29.6989
19.4726 | -35.1097 | -29.7058
19.4787 | -35.1101 | -29.7125
19.4848 | -35.1106 | -29.7196
19.4909 | -35.1111 | -29.7266
19.497 | -35.1116 | -29.7337
19.5031 | -35.112 | -29.7407
19.5092 | -35.1124 | -29.7477
19.5153 | -35.1129 | -29.7546
19.5214 | -35.1133 | -29.7612
19.5275 | -35.1138 | -29.7682
19.5336 | -35.1142 | -29.775
19.5397 | -35.1147 | -29.7819
19.5458 | -35.1152 | -29.7889
19.5519 | -35.1157 | -29.7958
19.5581 | -35.1161 | -29.8026
19.5642 | -35.1166 | -29.8094
19.5703 | -35.117 | -29.8163
19.5764 | -35.1175 | -29.8232
19.5825 | -35.1179 | -29.8302
19.5886 | -35.1184 | -29.8372
19.5947 | -35.1188 | -29.844
19.6008 | -35.1192 | -29.851
19.6069 | -35.1197 | -29.8579
19.613 | -35.1202 | -29.865
19.6191 | -35.1206 | -29.8718
19.6252 | -35.121 | -29.8788
19.6313 | -35.1215 | -29.8855
19.6374 | -35.1219 | -29.8926
19.6435 | -35.1224 | -29.8994
19.6496 | -35.1229 | -29.9063
19.6557 | -35.1233 | -29.9133
19.6618 | -35.1238 | -29.9197
19.6679 | -35.1242 | -29.9266
19.674 | -35.1247 | -29.9335
19.6801 | -35.1251 | -29.9402
19.6862 | -35.1256 | -29.9472
19.6923 | -35.126 | -29.9542
19.6984 | -35.1265 | -29.9605
19.7046 | -35.127 | -29.9674
19.7107 | -35.1274 | -29.9742
19.7168 | -35.1279 | -29.9812
19.7229 | -35.1283 | -29.9882
19.729 | -35.1288 | -29.9951
19.7351 | -35.1293 | -30.0019
19.7412 | -35.1297 | -30.0088
19.7473 | -35.1302 | -30.0157
19.7534 | -35.1307 | -30.0227
19.7595 | -35.1311 | -30.0298
19.7656 | -35.1316 | -30.0368
19.7717 | -35.132 | -30.0435
19.7778 | -35.1325 | -30.0505
19.7839 | -35.1329 | -30.0574
19.79 | -35.1334 | -30.0642
19.7961 | -35.1338 | -30.0712
19.8022 | -35.1343 | -30.078
19.8083 | -35.1348 | -30.085
19.8144 | -35.1352 | -30.0919
19.8205 | -35.1357 | -30.0988
19.8266 | -35.1361 | -30.1058
19.8327 | -35.1366 | -30.1127
19.8388 | -35.137 | -30.1192
19.845 | -35.1375 | -30.126
19.8511 | -35.1379 | -30.133
19.8572 | -35.1384 | -30.14
19.8633 | -35.1389 | -30.1467
19.8694 | -35.1393 | -30.1536
19.8755 | -35.1398 | -30.1605
19.8816 | -35.1403 | -30.1675
19.8877 | -35.1407 | -30.1744
19.8938 | -35.1412 | -30.1812
19.8999 | -35.1416 | -30.1883
19.906 | -35.1421 | -30.1952
19.9121 | -35.1426 | -30.2018
19.9182 | -35.143 | -30.2088
19.9243 | -35.1435 | -30.2158
19.9304 | -35.144 | -30.2229
19.9365 | -35.1444 | -30.2299
19.9426 | -35.1449 | -30.2365
19.9487 | -35.1453 | -30.2435
19.9548 | -35.1458 | -30.2504
19.9609 | -35.1463 | -30.257
19.967 | -35.1467 | -30.2639
19.9731 | -35.1472 | -30.2705
19.9792 | -35.1476 | -30.2774
19.9853 | -35.1481 | -30.2842
19.9915 | -35.1486 | -30.2911
19.9976 | -35.149 | -30.298
20.0037 | -35.1494 | -30.3048
20.0098 | -35.1498 | -30.3117
20.0159 | -35.1503 | -30.3185
20.022 | -35.1507 | -30.3253
20.0281 | -35.1512 | -30.3322
20.0342 | -35.1517 | -30.3389
20.0403 | -35.1522 | -30.3458
20.0464 | -35.1526 | -30.3525
20.0525 | -35.1531 | -30.359
20.0586 | -35.1536 | -30.3657
20.0647 | -35.154 | -30.3724
20.0708 | -35.1545 | -30.3794
20.0769 | -35.1549 | -30.3862
20.083 | -35.1553 | -30.3929
20.0891 | -35.1557 | -30.3997
20.0952 | -35.1562 | -30.4064
20.1013 | -35.1567 | -30.4128
20.1074 | -35.1571 | -30.4193
20.1135 | -35.1576 | -30.4259
20.1196 | -35.158 | -30.4323
20.1257 | -35.1585 | -30.4394
20.1319 | -35.159 | -30.4464
20.138 | -35.1594 | -30.4531
20.1441 | -35.1599 | -30.4601
20.1502 | -35.1604 | -30.4671
20.1563 | -35.1608 | -30.4739
20.1624 | -35.1613 | -30.4808
20.1685 | -35.1618 | -30.4877
20.1746 | -35.1622 | -30.4944
20.1807 | -35.1626 | -30.5011
20.1868 | -35.163 | -30.5078
20.1929 | -35.1635 | -30.5149
20.199 | -35.164 | -30.5219
20.2051 | -35.1644 | -30.5288
20.2112 | -35.1649 | -30.5359
20.2173 | -35.1654 | -30.5431
20.2234 | -35.1658 | -30.5497
20.2295 | -35.1663 | -30.5564
20.2356 | -35.1668 | -30.5635
20.2417 | -35.1673 | -30.5703
20.2478 | -35.1677 | -30.577
20.2539 | -35.1681 | -30.5838
20.26 | -35.1685 | -30.5902
20.2661 | -35.169 | -30.5967
20.2723 | -35.1694 | -30.6037
20.2784 | -35.1699 | -30.6099
20.2845 | -35.1704 | -30.6166
20.2906 | -35.1708 | -30.6232
20.2967 | -35.1712 | -30.6298
20.3028 | -35.1717 | -30.6364
20.3089 | -35.1721 | -30.6429
20.315 | -35.1726 | -30.6496
20.3211 | -35.173 | -30.6566
20.3272 | -35.1735 | -30.6635
20.3333 | -35.174 | -30.6699
20.3394 | -35.1744 | -30.6766
20.3455 | -35.1749 | -30.6835
20.3516 | -35.1753 | -30.6899
20.3577 | -35.1758 | -30.6965
20.3638 | -35.1762 | -30.7031
20.3699 | -35.1767 | -30.7096
20.376 | -35.1771 | -30.7161
20.3821 | -35.1776 | -30.7229
20.3882 | -35.1781 | -30.7296
20.3943 | -35.1785 | -30.7364
20.4004 | -35.179 | -30.7426
20.4065 | -35.1794 | -30.7491
20.4126 | -35.1799 | -30.7559
20.4188 | -35.1803 | -30.7626
20.4249 | -35.1807 | -30.7694
20.431 | -35.1811 | -30.7759
20.4371 | -35.1816 | -30.7823
20.4432 | -35.1821 | -30.7892
20.4493 | -35.1825 | -30.7959
20.4554 | -35.1829 | -30.8026
20.4615 | -35.1833 | -30.8094
20.4676 | -35.1838 | -30.8162
20.4737 | -35.1842 | -30.8225
20.4798 | -35.1847 | -30.8291
20.4859 | -35.1851 | -30.8355
20.492 | -35.1856 | -30.8422
20.4981 | -35.1861 | -30.8488
20.5042 | -35.1865 | -30.8556
20.5103 | -35.1869 | -30.8621
20.5164 | -35.1874 | -30.8686
20.5225 | -35.1878 | -30.8751
20.5286 | -35.1883 | -30.8818
20.5347 | -35.1887 | -30.8885
20.5408 | -35.1892 | -30.8947
20.5469 | -35.1896 | -30.9015
20.553 | -35.1901 | -30.908
20.5592 | -35.1905 | -30.9148
20.5653 | -35.191 | -30.9211
20.5714 | -35.1915 | -30.9279
20.5775 | -35.192 | -30.9346
20.5836 | -35.1924 | -30.9412
20.5897 | -35.1929 | -30.9481
20.5958 | -35.1933 | -30.9546
20.6019 | -35.1938 | -30.9614
20.608 | -35.1943 | -30.968
20.6141 | -35.1948 | -30.9744
20.6202 | -35.1952 | -30.9809
20.6263 | -35.1957 | -30.9876
20.6324 | -35.1962 | -30.9943
20.6385 | -35.1967 | -31.001
20.6446 | -35.1971 | -31.0077
20.6507 | -35.1975 | -31.0141
20.6568 | -35.198 | -31.0208
20.6629 | -35.1984 | -31.0274
20.669 | -35.1989 | -31.0343
20.6751 | -35.1993 | -31.0413
20.6812 | -35.1998 | -31.0476
20.6873 | -35.2002 | -31.0542
20.6934 | -35.2007 | -31.061
20.6995 | -35.2012 | -31.0677
20.7057 | -35.2016 | -31.0745
20.7118 | -35.2021 | -31.081
20.7179 | -35.2026 | -31.0875
20.724 | -35.203 | -31.094
20.7301 | -35.2035 | -31.1005
20.7362 | -35.2039 | -31.1069
20.7423 | -35.2044 | -31.1134
20.7484 | -35.2048 | -31.1199
20.7545 | -35.2053 | -31.1261
20.7606 | -35.2058 | -31.1329
20.7667 | -35.2063 | -31.1394
20.7728 | -35.2066 | -31.1461
20.7789 | -35.2071 | -31.1528
20.785 | -35.2075 | -31.1592
20.7911 | -35.208 | -31.1659
20.7972 | -35.2084 | -31.1723
20.8033 | -35.2089 | -31.179
20.8094 | -35.2093 | -31.1854
20.8155 | -35.2098 | -31.1918
20.8216 | -35.2102 | -31.1981
20.8277 | -35.2107 | -31.2039
20.8338 | -35.2111 | -31.2104
20.8399 | -35.2115 | -31.2167
20.8461 | -35.212 | -31.2234
20.8522 | -35.2124 | -31.23
20.8583 | -35.2129 | -31.2359
20.8644 | -35.2134 | -31.2426
20.8705 | -35.2138 | -31.2492
20.8766 | -35.2142 | -31.2557
20.8827 | -35.2147 | -31.2625
20.8888 | -35.2151 | -31.2692
20.8949 | -35.2156 | -31.2753
20.901 | -35.216 | -31.282
20.9071 | -35.2164 | -31.2886
20.9132 | -35.2169 | -31.2953
20.9193 | -35.2174 | -31.3019
20.9254 | -35.2178 | -31.3085
20.9315 | -35.2183 | -31.3156
20.9376 | -35.2187 | -31.3222
20.9437 | -35.2191 | -31.3287
20.9498 | -35.2196 | -31.3348
20.9559 | -35.22 | -31.3414
20.962 | -35.2204 | -31.3478
20.9681 | -35.2209 | -31.3547
20.9742 | -35.2214 | -31.3611
20.9803 | -35.2218 | -31.3675
20.9864 | -35.2224 | -31.3743
20.9926 | -35.2228 | -31.3813
20.9987 | -35.2233 | -31.3873
21.0048 | -35.2238 | -31.394
21.0109 | -35.2243 | -31.4005
21.017 | -35.2247 | -31.4072
21.0231 | -35.2252 | -31.4135
21.0292 | -35.2256 | -31.4201
21.0353 | -35.2261 | -31.4262
21.0414 | -35.2265 | -31.4328
21.0475 | -35.2269 | -31.4395
21.0536 | -35.2273 | -31.4459
21.0597 | -35.2278 | -31.4521
21.0658 | -35.2283 | -31.4587
21.0719 | -35.2288 | -31.4645
21.078 | -35.2291 | -31.4712
21.0841 | -35.2296 | -31.4777
21.0902 | -35.23 | -31.484
21.0963 | -35.2305 | -31.4904
21.1024 | -35.2309 | -31.4965
21.1085 | -35.2314 | -31.5032
21.1146 | -35.2318 | -31.5095
21.1207 | -35.2322 | -31.5162
21.1268 | -35.2327 | -31.5229
21.133 | -35.2332 | -31.5293
21.1391 | -35.2338 | -31.536
21.1452 | -35.2342 | -31.5427
21.1513 | -35.2346 | -31.5488
21.1574 | -35.2351 | -31.5555
21.1635 | -35.2356 | -31.5619
21.1696 | -35.236 | -31.5681
21.1757 | -35.2365 | -31.5739
21.1818 | -35.2369 | -31.5804
21.1879 | -35.2373 | -31.5873
21.194 | -35.2378 | -31.5937
21.2001 | -35.2382 | -31.6002
21.2062 | -35.2387 | -31.6066
21.2123 | -35.2391 | -31.6125
21.2184 | -35.2396 | -31.6194
21.2245 | -35.24 | -31.6256
21.2306 | -35.2404 | -31.6322
21.2367 | -35.2408 | -31.639
21.2428 | -35.2413 | -31.6447
21.2489 | -35.2418 | -31.651
21.255 | -35.2422 | -31.6576
21.2611 | -35.2427 | -31.6642
21.2672 | -35.2432 | -31.6708
21.2733 | -35.2436 | -31.6772
21.2795 | -35.244 | -31.6832
21.2856 | -35.2444 | -31.6893
21.2917 | -35.2449 | -31.6954
21.2978 | -35.2453 | -31.7017
21.3039 | -35.2458 | -31.7082
21.31 | -35.2463 | -31.7144
21.3161 | -35.2468 | -31.7206
21.3222 | -35.2473 | -31.7274
21.3283 | -35.2477 | -31.7343
21.3344 | -35.2482 | -31.7407
21.3405 | -35.2487 | -31.7467
21.3466 | -35.2492 | -31.7532
21.3527 | -35.2496 | -31.7594
21.3588 | -35.2501 | -31.7658
21.3649 | -35.2505 | -31.7721
21.371 | -35.251 | -31.7787
21.3771 | -35.2514 | -31.7849
21.3832 | -35.2519 | -31.7912
21.3893 | -35.2524 | -31.7975
21.3954 | -35.2528 | -31.8043
21.4015 | -35.2533 | -31.8104
21.4076 | -35.2537 | -31.817
21.4137 | -35.2541 | -31.8231
21.4199 | -35.2546 | -31.8297
21.426 | -35.2549 | -31.8354
21.4321 | -35.2554 | -31.8415
21.4382 | -35.2558 | -31.8483
21.4443 | -35.2563 | -31.8547
21.4504 | -35.2568 | -31.8611
21.4565 | -35.2572 | -31.8667
21.4626 | -35.2576 | -31.8729
21.4687 | -35.2581 | -31.8796
21.4748 | -35.2585 | -31.886
21.4809 | -35.2589 | -31.8926
21.487 | -35.2593 | -31.8985
21.4931 | -35.2596 | -31.9051
21.4992 | -35.2601 | -31.9113
21.5053 | -35.2605 | -31.9176
21.5114 | -35.261 | -31.9241
21.5175 | -35.2614 | -31.9303
21.5236 | -35.2619 | -31.9368
21.5297 | -35.2623 | -31.9435
21.5358 | -35.2628 | -31.9499
21.5419 | -35.2632 | -31.956
21.548 | -35.2636 | -31.9623
21.5541 | -35.2641 | -31.9687
21.5602 | -35.2645 | -31.9749
21.5664 | -35.2651 | -31.981
21.5725 | -35.2655 | -31.9876
21.5786 | -35.266 | -31.9938
21.5847 | -35.2665 | -32.0006
21.5908 | -35.267 | -32.0066
21.5969 | -35.2674 | -32.0123
21.603 | -35.2679 | -32.0187
21.6091 | -35.2683 | -32.025
21.6152 | -35.2687 | -32.0315
21.6213 | -35.2691 | -32.0379
21.6274 | -35.2696 | -32.0444
21.6335 | -35.27 | -32.0506
21.6396 | -35.2704 | -32.057
21.6457 | -35.2708 | -32.0636
21.6518 | -35.2712 | -32.0699
21.6579 | -35.2717 | -32.0761
21.664 | -35.2723 | -32.0823
21.6701 | -35.2727 | -32.0882
21.6762 | -35.2731 | -32.0946
21.6823 | -35.2736 | -32.1008
21.6884 | -35.274 | -32.1072
21.6945 | -35.2745 | -32.1139
21.7006 | -35.2749 | -32.1204
21.7068 | -35.2754 | -32.1259
21.7129 | -35.2758 | -32.132
21.719 | -35.2763 | -32.1381
21.7251 | -35.2767 | -32.1444
21.7312 | -35.2771 | -32.1507
21.7373 | -35.2775 | -32.1561
21.7434 | -35.278 | -32.1627
21.7495 | -35.2785 | -32.1689
21.7556 | -35.2789 | -32.1754
21.7617 | -35.2793 | -32.1818
21.7678 | -35.2797 | -32.188
21.7739 | -35.2801 | -32.1937
21.78 | -35.2805 | -32.1998
21.7861 | -35.2809 | -32.2064
21.7922 | -35.2813 | -32.2124
21.7983 | -35.2817 | -32.2185
21.8044 | -35.2822 | -32.225
21.8105 | -35.2827 | -32.2312
21.8166 | -35.2831 | -32.2379
21.8227 | -35.2836 | -32.2442
21.8288 | -35.284 | -32.251
21.8349 | -35.2844 | -32.2569
21.841 | -35.2849 | -32.2629
21.8471 | -35.2853 | -32.2688
21.8533 | -35.2858 | -32.2749
21.8594 | -35.2863 | -32.281
21.8655 | -35.2867 | -32.2874
21.8716 | -35.2871 | -32.2936
21.8777 | -35.2876 | -32.3001
21.8838 | -35.2881 | -32.3063
21.8899 | -35.2885 | -32.3124
21.896 | -35.2889 | -32.3186
21.9021 | -35.2894 | -32.3247
21.9082 | -35.2898 | -32.3309
21.9143 | -35.2903 | -32.3372
21.9204 | -35.2906 | -32.3432
21.9265 | -35.291 | -32.3493
21.9326 | -35.2915 | -32.3558
21.9387 | -35.2919 | -32.3625
21.9448 | -35.2924 | -32.3687
21.9509 | -35.293 | -32.3752
21.957 | -35.2934 | -32.3816
21.9631 | -35.2938 | -32.3878
21.9692 | -35.2942 | -32.3938
21.9753 | -35.2946 | -32.3995
21.9814 | -35.295 | -32.4059
21.9875 | -35.2955 | -32.4115
21.9937 | -35.2959 | -32.4177
21.9998 | -35.2964 | -32.4235
22.0059 | -35.2968 | -32.4295
22.012 | -35.2973 | -32.4354
22.0181 | -35.2977 | -32.4413
22.0242 | -35.2981 | -32.4475
22.0303 | -35.2986 | -32.4538
22.0364 | -35.299 | -32.46
22.0425 | -35.2995 | -32.4661
22.0486 | -35.2999 | -32.4721
22.0547 | -35.3003 | -32.4776
22.0608 | -35.3007 | -32.4836
22.0669 | -35.3012 | -32.49
22.073 | -35.3017 | -32.4961
22.0791 | -35.3022 | -32.5024
22.0852 | -35.3026 | -32.5087
22.0913 | -35.3029 | -32.5146
22.0974 | -35.3033 | -32.5203
22.1035 | -35.3038 | -32.5263
22.1096 | -35.3043 | -32.5328
22.1157 | -35.3048 | -32.5393
22.1218 | -35.3052 | -32.5453
22.1279 | -35.3057 | -32.5512
22.134 | -35.3062 | -32.5574
22.1402 | -35.3066 | -32.5636
22.1463 | -35.3069 | -32.5697
22.1524 | -35.3074 | -32.5758
22.1585 | -35.3078 | -32.5819
22.1646 | -35.3082 | -32.5882
22.1707 | -35.3087 | -32.5943
22.1768 | -35.3092 | -32.6006
22.1829 | -35.3096 | -32.6067
22.189 | -35.3101 | -32.6129
22.1951 | -35.3105 | -32.6189
22.2012 | -35.3109 | -32.6248
22.2073 | -35.3114 | -32.6313
22.2134 | -35.3119 | -32.638
22.2195 | -35.3123 | -32.6442
22.2256 | -35.3127 | -32.6504
22.2317 | -35.313 | -32.6566
22.2378 | -35.3134 | -32.6627
22.2439 | -35.3139 | -32.6691
22.25 | -35.3143 | -32.6753
22.2561 | -35.3148 | -32.6815
22.2622 | -35.3152 | -32.6875
22.2683 | -35.3157 | -32.6933
22.2744 | -35.3161 | -32.699
22.2806 | -35.3165 | -32.7059
22.2867 | -35.3169 | -32.7118
22.2928 | -35.3173 | -32.7179
22.2989 | -35.3178 | -32.7241
22.305 | -35.3183 | -32.7306
22.3111 | -35.3187 | -32.7366
22.3172 | -35.3192 | -32.7428
22.3233 | -35.3196 | -32.7484
22.3294 | -35.3201 | -32.7549
22.3355 | -35.3205 | -32.7611
22.3416 | -35.3209 | -32.7669
22.3477 | -35.3213 | -32.7731
22.3538 | -35.3218 | -32.7793
22.3599 | -35.3222 | -32.7855
22.366 | -35.3227 | -32.7915
22.3721 | -35.3232 | -32.7972
22.3782 | -35.3236 | -32.8029
22.3843 | -35.324 | -32.809
22.3904 | -35.3245 | -32.8155
22.3965 | -35.3249 | -32.8217
22.4026 | -35.3253 | -32.8277
22.4087 | -35.3257 | -32.8326
22.4148 | -35.3261 | -32.8391
22.4209 | -35.3265 | -32.8449
22.4271 | -35.327 | -32.851
22.4332 | -35.3274 | -32.8571
22.4393 | -35.3279 | -32.863
22.4454 | -35.3284 | -32.8691
22.4515 | -35.3289 | -32.8749
22.4576 | -35.3293 | -32.881
22.4637 | -35.3298 | -32.8876
22.4698 | -35.3302 | -32.8936
22.4759 | -35.3308 | -32.8994
22.482 | -35.3312 | -32.9054
22.4881 | -35.3316 | -32.9119
22.4942 | -35.3321 | -32.9183
22.5003 | -35.3324 | -32.9244
22.5064 | -35.3329 | -32.9308
22.5125 | -35.3332 | -32.9366
22.5186 | -35.3337 | -32.9426
22.5247 | -35.3341 | -32.949
22.5308 | -35.3345 | -32.9552
22.5369 | -35.335 | -32.9611
22.543 | -35.3353 | -32.9672
22.5491 | -35.3357 | -32.9726
22.5552 | -35.3362 | -32.9786
22.5613 | -35.3366 | -32.9848
22.5675 | -35.337 | -32.9905
22.5736 | -35.3374 | -32.9968
22.5797 | -35.3379 | -33.0024
22.5858 | -35.3383 | -33.0079
22.5919 | -35.3387 | -33.0139
22.598 | -35.3392 | -33.0202
22.6041 | -35.3395 | -33.0261
22.6102 | -35.3399 | -33.0325
22.6163 | -35.3403 | -33.0385
22.6224 | -35.3407 | -33.045
22.6285 | -35.3412 | -33.0507
22.6346 | -35.3416 | -33.0566
22.6407 | -35.3421 | -33.0623
22.6468 | -35.3425 | -33.0681
22.6529 | -35.3429 | -33.0739
22.659 | -35.3434 | -33.0797
22.6651 | -35.3438 | -33.0854
22.6712 | -35.3441 | -33.0913
22.6773 | -35.3445 | -33.0976
22.6834 | -35.3449 | -33.1036
22.6895 | -35.3454 | -33.1092
22.6956 | -35.3458 | -33.1154
22.7017 | -35.3463 | -33.1212
22.7079 | -35.3467 | -33.127
22.714 | -35.3471 | -33.133
22.7201 | -35.3476 | -33.1394
22.7262 | -35.348 | -33.1454
22.7323 | -35.3484 | -33.1516
22.7384 | -35.3488 | -33.1574
22.7445 | -35.3493 | -33.1634
22.7506 | -35.3497 | -33.1692
22.7567 | -35.3501 | -33.1749
22.7628 | -35.3505 | -33.1806
22.7689 | -35.351 | -33.1866
22.775 | -35.3514 | -33.1923
22.7811 | -35.3518 | -33.1983
22.7872 | -35.3522 | -33.2045
22.7933 | -35.3527 | -33.2106
22.7994 | -35.3531 | -33.2166
22.8055 | -35.3535 | -33.2228
22.8116 | -35.354 | -33.2286
22.8177 | -35.3545 | -33.235
22.8238 | -35.3549 | -33.2405
22.8299 | -35.3554 | -33.2459
22.836 | -35.3558 | -33.2517
22.8421 | -35.3562 | -33.2577
22.8482 | -35.3566 | -33.2637
22.8544 | -35.3569 | -33.2696
22.8605 | -35.3574 | -33.2756
22.8666 | -35.3578 | -33.2813
22.8727 | -35.3583 | -33.2873
22.8788 | -35.3587 | -33.2931
22.8849 | -35.3591 | -33.299
22.891 | -35.3596 | -33.3052
22.8971 | -35.36 | -33.3113
22.9032 | -35.3604 | -33.317
22.9093 | -35.3609 | -33.3233
22.9154 | -35.3613 | -33.3293
22.9215 | -35.3617 | -33.3353
22.9276 | -35.3621 | -33.3412
22.9337 | -35.3626 | -33.3467
22.9398 | -35.363 | -33.353
22.9459 | -35.3634 | -33.3594
22.952 | -35.3639 | -33.3655
22.9581 | -35.3644 | -33.3716
22.9642 | -35.3648 | -33.3773
22.9703 | -35.3652 | -33.3831
22.9764 | -35.3656 | -33.389
22.9825 | -35.3659 | -33.3952
22.9886 | -35.3664 | -33.4011
22.9948 | -35.3668 | -33.4071
23.0009 | -35.3672 | -33.4131
23.007 | -35.3676 | -33.419
23.0131 | -35.368 | -33.4253
23.0192 | -35.3684 | -33.4311
23.0253 | -35.3688 | -33.4372
23.0314 | -35.3691 | -33.4433
23.0375 | -35.3696 | -33.4494
23.0436 | -35.37 | -33.4544
23.0497 | -35.3705 | -33.4602
23.0558 | -35.3709 | -33.4657
23.0619 | -35.3714 | -33.4715
23.068 | -35.3718 | -33.4776
23.0741 | -35.3722 | -33.4831
23.0802 | -35.3726 | -33.489
23.0863 | -35.373 | -33.4945
23.0924 | -35.3734 | -33.5001
23.0985 | -35.3738 | -33.5061
23.1046 | -35.3742 | -33.512
23.1107 | -35.3746 | -33.5184
23.1168 | -35.3751 | -33.524
23.1229 | -35.3755 | -33.5297
23.129 | -35.3759 | -33.5353
23.1351 | -35.3763 | -33.541
23.1413 | -35.3767 | -33.5468
23.1474 | -35.3772 | -33.5524
23.1535 | -35.3776 | -33.5581
23.1596 | -35.378 | -33.5638
23.1657 | -35.3785 | -33.5698
23.1718 | -35.3789 | -33.5758
23.1779 | -35.3794 | -33.5815
23.184 | -35.3799 | -33.5868
23.1901 | -35.3803 | -33.5927
23.1962 | -35.3807 | -33.5988
23.2023 | -35.3812 | -33.6048
23.2084 | -35.3816 | -33.6108
23.2145 | -35.382 | -33.6166
23.2206 | -35.3824 | -33.6226
23.2267 | -35.3828 | -33.6283
23.2328 | -35.3832 | -33.6338
23.2389 | -35.3836 | -33.6395
23.245 | -35.384 | -33.6453
23.2511 | -35.3844 | -33.6509
23.2572 | -35.3848 | -33.6571
23.2633 | -35.3852 | -33.6629
23.2694 | -35.3857 | -33.6688
23.2755 | -35.3861 | -33.6744
23.2817 | -35.3865 | -33.6802
23.2878 | -35.3869 | -33.6857
23.2939 | -35.3873 | -33.6918
23.3 | -35.3878 | -33.6976
23.3061 | -35.3881 | -33.7035
23.3122 | -35.3886 | -33.7093
23.3183 | -35.389 | -33.7151
23.3244 | -35.3894 | -33.7211
23.3305 | -35.3898 | -33.7273
23.3366 | -35.3903 | -33.7333
23.3427 | -35.3907 | -33.7389
23.3488 | -35.3911 | -33.7448
23.3549 | -35.3915 | -33.7499
23.361 | -35.3919 | -33.7553
23.3671 | -35.3924 | -33.7612
23.3732 | -35.3928 | -33.767
23.3793 | -35.3932 | -33.773
23.3854 | -35.3936 | -33.7792
23.3915 | -35.3941 | -33.7846
23.3976 | -35.3945 | -33.7906
23.4037 | -35.3949 | -33.7964
23.4098 | -35.3953 | -33.8025
23.4159 | -35.3958 | -33.8082
23.422 | -35.3962 | -33.8137
23.4282 | -35.3967 | -33.8194
23.4343 | -35.3971 | -33.8253
23.4404 | -35.3975 | -33.8311
23.4465 | -35.3979 | -33.8368
23.4526 | -35.3984 | -33.8428
23.4587 | -35.3988 | -33.8484
23.4648 | -35.3993 | -33.8538
23.4709 | -35.3997 | -33.8598
23.477 | -35.4001 | -33.8658
23.4831 | -35.4005 | -33.8716
23.4892 | -35.4009 | -33.8775
23.4953 | -35.4013 | -33.8833
23.5014 | -35.4018 | -33.8893
23.5075 | -35.4023 | -33.895
23.5136 | -35.4027 | -33.9007
23.5197 | -35.4031 | -33.9063
23.5258 | -35.4034 | -33.9118
23.5319 | -35.4039 | -33.9171
23.538 | -35.4043 | -33.923
23.5441 | -35.4046 | -33.9289
23.5502 | -35.4051 | -33.9346
23.5563 | -35.4056 | -33.9403
23.5624 | -35.406 | -33.9461
23.5686 | -35.4064 | -33.952
23.5747 | -35.4068 | -33.9578
23.5808 | -35.4072 | -33.9633
23.5869 | -35.4077 | -33.9688
23.593 | -35.4081 | -33.9745
23.5991 | -35.4085 | -33.9803
23.6052 | -35.4089 | -33.986
23.6113 | -35.4093 | -33.9922
23.6174 | -35.4097 | -33.9984
23.6235 | -35.4102 | -34.0041
23.6296 | -35.4106 | -34.01
23.6357 | -35.411 | -34.0155
23.6418 | -35.4114 | -34.0207
23.6479 | -35.4118 | -34.0264
23.654 | -35.4122 | -34.032
23.6601 | -35.4127 | -34.0379
23.6662 | -35.4131 | -34.0436
23.6723 | -35.4135 | -34.0498
23.6784 | -35.4139 | -34.0554
23.6845 | -35.4142 | -34.0611
23.6906 | -35.4147 | -34.0665
23.6967 | -35.4151 | -34.0725
23.7028 | -35.4154 | -34.0785
23.7089 | -35.4158 | -34.084
23.7151 | -35.4162 | -34.0899
23.7212 | -35.4167 | -34.0957
23.7273 | -35.4171 | -34.1016
23.7334 | -35.4175 | -34.1068
23.7395 | -35.4179 | -34.1125
23.7456 | -35.4183 | -34.1178
23.7517 | -35.4187 | -34.1234
23.7578 | -35.4191 | -34.129
23.7639 | -35.4195 | -34.135
23.77 | -35.42 | -34.1407
23.7761 | -35.4203 | -34.1468
23.7822 | -35.4208 | -34.1526
23.7883 | -35.4212 | -34.1581
23.7944 | -35.4216 | -34.1641
23.8005 | -35.422 | -34.1699
23.8066 | -35.4224 | -34.1759
23.8127 | -35.4229 | -34.1817
23.8188 | -35.4232 | -34.1878
23.8249 | -35.4236 | -34.1938
23.831 | -35.424 | -34.1994
23.8371 | -35.4244 | -34.2053
23.8432 | -35.4248 | -34.2111
23.8493 | -35.4253 | -34.2167
23.8555 | -35.4257 | -34.2225
23.8616 | -35.4261 | -34.2285
23.8677 | -35.4265 | -34.2345
23.8738 | -35.4269 | -34.2404
23.8799 | -35.4273 | -34.2463
23.886 | -35.4277 | -34.2519
23.8921 | -35.4281 | -34.2578
23.8982 | -35.4286 | -34.2636
23.9043 | -35.429 | -34.2693
23.9104 | -35.4294 | -34.2747
23.9165 | -35.4298 | -34.2807
23.9226 | -35.4302 | -34.2862
23.9287 | -35.4306 | -34.2923
23.9348 | -35.431 | -34.2981
23.9409 | -35.4315 | -34.3036
23.947 | -35.4319 | -34.3094
23.9531 | -35.4322 | -34.3147
23.9592 | -35.4327 | -34.3207
23.9653 | -35.433 | -34.3265
23.9714 | -35.4334 | -34.3323
23.9775 | -35.4338 | -34.338
23.9836 | -35.4343 | -34.3436
23.9897 | -35.4347 | -34.3495
23.9958 | -35.4351 | -34.3549
24.002 | -35.4355 | -34.3607
24.0081 | -35.4359 | -34.3658
24.0142 | -35.4363 | -34.3719
24.0203 | -35.4367 | -34.3778
24.0264 | -35.4372 | -34.3837
24.0325 | -35.4376 | -34.3895
24.0386 | -35.438 | -34.3955
24.0447 | -35.4385 | -34.401
24.0508 | -35.4388 | -34.4068
24.0569 | -35.4393 | -34.4125
24.063 | -35.4398 | -34.4182
24.0691 | -35.4402 | -34.4239
24.0752 | -35.4406 | -34.4296
24.0813 | -35.4411 | -34.4352
24.0874 | -35.4415 | -34.4406
24.0935 | -35.4419 | -34.4466
24.0996 | -35.4423 | -34.4523
24.1057 | -35.4428 | -34.4579
24.1118 | -35.4432 | -34.464
24.1179 | -35.4436 | -34.4702
24.124 | -35.444 | -34.4757
24.1301 | -35.4444 | -34.481
24.1362 | -35.4448 | -34.4871
24.1424 | -35.4452 | -34.493
24.1485 | -35.4456 | -34.4985
24.1546 | -35.446 | -34.5039
24.1607 | -35.4463 | -34.5096
24.1668 | -35.4467 | -34.5151
24.1729 | -35.4471 | -34.5208
24.179 | -35.4475 | -34.5262
24.1851 | -35.448 | -34.5319
24.1912 | -35.4483 | -34.5374
24.1973 | -35.4487 | -34.5431
24.2034 | -35.4491 | -34.5494
24.2095 | -35.4496 | -34.555
24.2156 | -35.45 | -34.5607
24.2217 | -35.4504 | -34.5661
24.2278 | -35.4508 | -34.5716
24.2339 | -35.4512 | -34.5778
24.24 | -35.4516 | -34.5835
24.2461 | -35.452 | -34.5889
24.2522 | -35.4524 | -34.5941
24.2583 | -35.4528 | -34.5999
24.2644 | -35.4532 | -34.6054
24.2705 | -35.4536 | -34.6113
24.2766 | -35.4541 | -34.6168
24.2827 | -35.4545 | -34.6224
24.2889 | -35.4549 | -34.6284
24.295 | -35.4553 | -34.6342
24.3011 | -35.4557 | -34.6395
24.3072 | -35.4562 | -34.6452
24.3133 | -35.4566 | -34.6506
24.3194 | -35.457 | -34.6563
24.3255 | -35.4574 | -34.6619
24.3316 | -35.4579 | -34.6677
24.3377 | -35.4582 | -34.6732
24.3438 | -35.4587 | -34.6791
24.3499 | -35.459 | -34.6846
24.356 | -35.4594 | -34.6902
24.3621 | -35.4599 | -34.696
24.3682 | -35.4602 | -34.7014
24.3743 | -35.4606 | -34.7069
24.3804 | -35.4609 | -34.7127
24.3865 | -35.4613 | -34.7183
24.3926 | -35.4616 | -34.724
24.3987 | -35.4621 | -34.7296
24.4048 | -35.4625 | -34.735
24.4109 | -35.4629 | -34.7403
24.417 | -35.4633 | -34.7463
24.4231 | -35.4637 | -34.7522
24.4293 | -35.4641 | -34.7576
24.4354 | -35.4645 | -34.7636
24.4415 | -35.4649 | -34.7688
24.4476 | -35.4653 | -34.7748
24.4537 | -35.4657 | -34.7804
24.4598 | -35.4661 | -34.7857
24.4659 | -35.4666 | -34.7916
24.472 | -35.467 | -34.7971
24.4781 | -35.4674 | -34.8025
24.4842 | -35.4677 | -34.8089
24.4903 | -35.4681 | -34.8143
24.4964 | -35.4685 | -34.8203
24.5025 | -35.4689 | -34.8258
24.5086 | -35.4693 | -34.8312
24.5147 | -35.4697 | -34.8371
24.5208 | -35.4701 | -34.8423
24.5269 | -35.4706 | -34.8481
24.533 | -35.471 | -34.8539
24.5391 | -35.4715 | -34.8597
24.5452 | -35.4719 | -34.8655
24.5513 | -35.4723 | -34.8713
24.5574 | -35.4726 | -34.8768
24.5635 | -35.473 | -34.8824
24.5696 | -35.4734 | -34.888
24.5758 | -35.4738 | -34.8938
24.5819 | -35.4742 | -34.8998
24.588 | -35.4747 | -34.9057
24.5941 | -35.4751 | -34.9113
24.6002 | -35.4756 | -34.917
24.6063 | -35.4759 | -34.9232
24.6124 | -35.4763 | -34.929
24.6185 | -35.4767 | -34.935
24.6246 | -35.4771 | -34.9404
24.6307 | -35.4775 | -34.946
24.6368 | -35.4779 | -34.9518
24.6429 | -35.4783 | -34.9572
24.649 | -35.4787 | -34.9629
24.6551 | -35.4791 | -34.9684
24.6612 | -35.4795 | -34.9738
24.6673 | -35.4799 | -34.9794
24.6734 | -35.4803 | -34.9851
24.6795 | -35.4807 | -34.9908
24.6856 | -35.4811 | -34.9962
24.6917 | -35.4815 | -35.0015
24.6978 | -35.4818 | -35.0069
24.7039 | -35.4822 | -35.0125
24.71 | -35.4826 | -35.018
24.7162 | -35.483 | -35.0236
24.7223 | -35.4834 | -35.0292
24.7284 | -35.4838 | -35.0341
24.7345 | -35.4841 | -35.0396
24.7406 | -35.4846 | -35.0456
24.7467 | -35.485 | -35.0514
24.7528 | -35.4854 | -35.0566
24.7589 | -35.4858 | -35.0622
24.765 | -35.4862 | -35.0678
24.7711 | -35.4866 | -35.0738
24.7772 | -35.4871 | -35.0792
24.7833 | -35.4874 | -35.0846
24.7894 | -35.4879 | -35.0905
24.7955 | -35.4883 | -35.0962
24.8016 | -35.4887 | -35.1019
24.8077 | -35.4891 | -35.1079
24.8138 | -35.4895 | -35.1133
24.8199 | -35.4899 | -35.1194
24.826 | -35.4903 | -35.1249
24.8321 | -35.4907 | -35.1305
24.8382 | -35.4911 | -35.1353
24.8443 | -35.4915 | -35.1413
24.8504 | -35.492 | -35.147
24.8565 | -35.4924 | -35.1529
24.8627 | -35.4928 | -35.1583
24.8688 | -35.4931 | -35.1632
24.8749 | -35.4935 | -35.1686
24.881 | -35.4939 | -35.1741
24.8871 | -35.4943 | -35.1799
24.8932 | -35.4947 | -35.1856
24.8993 | -35.4952 | -35.1912
24.9054 | -35.4956 | -35.1971
24.9115 | -35.4961 | -35.2029
24.9176 | -35.4965 | -35.2087
24.9237 | -35.4968 | -35.2144
24.9298 | -35.4973 | -35.2203
24.9359 | -35.4977 | -35.2265
24.942 | -35.4981 | -35.2319
24.9481 | -35.4985 | -35.2376
24.9542 | -35.4989 | -35.2435
24.9603 | -35.4993 | -35.2489
24.9664 | -35.4997 | -35.2543
24.9725 | -35.5001 | -35.2599
24.9786 | -35.5005 | -35.2651
24.9847 | -35.5009 | -35.2709
24.9908 | -35.5013 | -35.2767
24.9969 | -35.5016 | -35.2825
25.0031 | -35.502 | -35.2879
25.0092 | -35.5023 | -35.293
25.0153 | -35.5026 | -35.2986
25.0214 | -35.503 | -35.3045
25.0275 | -35.5034 | -35.3102
25.0336 | -35.5038 | -35.3155
25.0397 | -35.5042 | -35.321
25.0458 | -35.5046 | -35.3267
25.0519 | -35.505 | -35.3328
25.058 | -35.5054 | -35.3384
25.0641 | -35.5058 | -35.3444
25.0702 | -35.5062 | -35.3499
25.0763 | -35.5066 | -35.3554
25.0824 | -35.507 | -35.3609
25.0885 | -35.5075 | -35.3667
25.0946 | -35.5079 | -35.3726
25.1007 | -35.5083 | -35.378
25.1068 | -35.5088 | -35.3834
25.1129 | -35.5093 | -35.3891
25.119 | -35.5098 | -35.3945
25.1251 | -35.5102 | -35.4001
25.1312 | -35.5105 | -35.4058
25.1373 | -35.511 | -35.4116
25.1435 | -35.5114 | -35.417
25.1496 | -35.5118 | -35.4225
25.1557 | -35.5122 | -35.4284
25.1618 | -35.5126 | -35.4341
25.1679 | -35.5131 | -35.4404
25.174 | -35.5135 | -35.4461
25.1801 | -35.5139 | -35.452
25.1862 | -35.5141 | -35.4576
25.1923 | -35.5146 | -35.4631
25.1984 | -35.515 | -35.4687
25.2045 | -35.5154 | -35.4747
25.2106 | -35.5158 | -35.4803
25.2167 | -35.5161 | -35.486
25.2228 | -35.5166 | -35.4913
25.2289 | -35.517 | -35.4964
25.235 | -35.5174 | -35.5023
25.2411 | -35.5179 | -35.5081
25.2472 | -35.5182 | -35.5141
25.2533 | -35.5186 | -35.5199
25.2594 | -35.5189 | -35.5249
25.2655 | -35.5193 | -35.5304
25.2716 | -35.5198 | -35.5358
25.2777 | -35.5201 | -35.5411
25.2838 | -35.5205 | -35.5468
25.29 | -35.5209 | -35.5525
25.2961 | -35.5213 | -35.5576
25.3022 | -35.5217 | -35.563
25.3083 | -35.5221 | -35.5686
25.3144 | -35.5226 | -35.5744
25.3205 | -35.523 | -35.58
25.3266 | -35.5234 | -35.5858
25.3327 | -35.5238 | -35.5914
25.3388 | -35.5241 | -35.5971
25.3449 | -35.5246 | -35.603
25.351 | -35.525 | -35.6082
25.3571 | -35.5254 | -35.6131
25.3632 | -35.5258 | -35.6189
25.3693 | -35.5262 | -35.6245
25.3754 | -35.5267 | -35.6303
25.3815 | -35.5271 | -35.636
25.3876 | -35.5275 | -35.6415
25.3937 | -35.5279 | -35.647
25.3998 | -35.5282 | -35.6522
25.4059 | -35.5286 | -35.6581
25.412 | -35.5291 | -35.6635
25.4181 | -35.5295 | -35.6685
25.4242 | -35.5298 | -35.6739
25.4304 | -35.5302 | -35.6795
25.4365 | -35.5306 | -35.6848
25.4426 | -35.531 | -35.6904
25.4487 | -35.5314 | -35.6965
25.4548 | -35.5317 | -35.7017
25.4609 | -35.5321 | -35.7075
25.467 | -35.5324 | -35.7127
25.4731 | -35.5329 | -35.7182
25.4792 | -35.5333 | -35.724
25.4853 | -35.5337 | -35.7298
25.4914 | -35.5341 | -35.7351
25.4975 | -35.5344 | -35.7407
25.5036 | -35.5348 | -35.7469
25.5097 | -35.5351 | -35.7526
25.5158 | -35.5355 | -35.7582
25.5219 | -35.536 | -35.7638
25.528 | -35.5364 | -35.7696
25.5341 | -35.5367 | -35.7751
25.5402 | -35.5371 | -35.7813
25.5463 | -35.5375 | -35.7866
25.5524 | -35.5379 | -35.7923
25.5585 | -35.5383 | -35.7981
25.5646 | -35.5387 | -35.804
25.5707 | -35.5391 | -35.8101
25.5769 | -35.5395 | -35.8154
25.583 | -35.54 | -35.8213
25.5891 | -35.5404 | -35.8269
25.5952 | -35.5408 | -35.8325
25.6013 | -35.5412 | -35.8375
25.6074 | -35.5416 | -35.8428
25.6135 | -35.5421 | -35.8484
25.6196 | -35.5425 | -35.8536
25.6257 | -35.5429 | -35.8591
25.6318 | -35.5433 | -35.8645
25.6379 | -35.5438 | -35.8699
25.644 | -35.5441 | -35.8761
25.6501 | -35.5445 | -35.8815
25.6562 | -35.545 | -35.887
25.6623 | -35.5454 | -35.8924
25.6684 | -35.5457 | -35.8979
25.6745 | -35.5461 | -35.9033
25.6806 | -35.5465 | -35.9089
25.6867 | -35.5468 | -35.9145
25.6928 | -35.5472 | -35.92
25.6989 | -35.5476 | -35.9252
25.705 | -35.548 | -35.9308
25.7111 | -35.5485 | -35.9367
25.7173 | -35.5489 | -35.9423
25.7234 | -35.5494 | -35.9481
25.7295 | -35.5497 | -35.9535
25.7356 | -35.5501 | -35.9594
25.7417 | -35.5505 | -35.965
25.7478 | -35.5508 | -35.9701
25.7539 | -35.5511 | -35.9757
25.76 | -35.5515 | -35.9811
25.7661 | -35.5519 | -35.9864
25.7722 | -35.5523 | -35.9922
25.7783 | -35.5527 | -35.9978
25.7844 | -35.5531 | -36.0032
25.7905 | -35.5534 | -36.0089
25.7966 | -35.5539 | -36.0147
25.8027 | -35.5543 | -36.0207
25.8088 | -35.5547 | -36.0264
25.8149 | -35.5551 | -36.0322
25.821 | -35.5554 | -36.0377
25.8271 | -35.5558 | -36.0434
25.8332 | -35.5563 | -36.0488
25.8393 | -35.5567 | -36.0545
25.8454 | -35.557 | -36.0603
25.8515 | -35.5573 | -36.0662
25.8576 | -35.5577 | -36.072
25.8638 | -35.5581 | -36.0779
25.8699 | -35.5584 | -36.0835
25.876 | -35.5588 | -36.0889
25.8821 | -35.5592 | -36.0943
25.8882 | -35.5595 | -36.1003
25.8943 | -35.5599 | -36.1063
25.9004 | -35.5603 | -36.1118
25.9065 | -35.5606 | -36.1178
25.9126 | -35.5611 | -36.1238
25.9187 | -35.5614 | -36.1297
25.9248 | -35.5619 | -36.1355
25.9309 | -35.5623 | -36.1412
25.937 | -35.5627 | -36.1467
25.9431 | -35.5631 | -36.1523
25.9492 | -35.5634 | -36.1583
25.9553 | -35.5638 | -36.1644
25.9614 | -35.5642 | -36.1697
25.9675 | -35.5646 | -36.1751
25.9736 | -35.565 | -36.1807
25.9797 | -35.5654 | -36.1858
25.9858 | -35.5658 | -36.1913
25.9919 | -35.5662 | -36.1971
25.998 | -35.5666 | -36.2029
26.0042 | -35.567 | -36.2084
26.0103 | -35.5674 | -36.2141
26.0164 | -35.5678 | -36.2197
26.0225 | -35.5681 | -36.2253
26.0286 | -35.5686 | -36.2308
26.0347 | -35.569 | -36.2368
26.0408 | -35.5695 | -36.2425
26.0469 | -35.5698 | -36.2477
26.053 | -35.5702 | -36.2535
26.0591 | -35.5706 | -36.2598
26.0652 | -35.5709 | -36.2657
26.0713 | -35.5712 | -36.2714
26.0774 | -35.5716 | -36.277
26.0835 | -35.572 | -36.2825
26.0896 | -35.5724 | -36.2886
26.0957 | -35.5728 | -36.294
26.1018 | -35.5733 | -36.3001
26.1079 | -35.5737 | -36.3054
26.114 | -35.574 | -36.3108
26.1201 | -35.5745 | -36.3161
26.1262 | -35.5749 | -36.3217
26.1323 | -35.5752 | -36.3274
26.1384 | -35.5756 | -36.3331
26.1445 | -35.576 | -36.3387
26.1507 | -35.5764 | -36.3445
26.1568 | -35.5768 | -36.3502
26.1629 | -35.5772 | -36.356
26.169 | -35.5776 | -36.3615
26.1751 | -35.5779 | -36.3666
26.1812 | -35.5782 | -36.3723
26.1873 | -35.5785 | -36.378
26.1934 | -35.5789 | -36.3838
26.1995 | -35.5793 | -36.3896
26.2056 | -35.5797 | -36.3956
26.2117 | -35.5801 | -36.401
26.2178 | -35.5805 | -36.4061
26.2239 | -35.5809 | -36.4121
26.23 | -35.5813 | -36.4177
26.2361 | -35.5817 | -36.4238
26.2422 | -35.5821 | -36.4297
26.2483 | -35.5825 | -36.435
26.2544 | -35.5829 | -36.4408
26.2605 | -35.5833 | -36.4463
26.2666 | -35.5838 | -36.4517
26.2727 | -35.5841 | -36.4575
26.2788 | -35.5845 | -36.4632
26.2849 | -35.5849 | -36.4687
26.2911 | -35.5853 | -36.474
26.2972 | -35.5857 | -36.4797
26.3033 | -35.5861 | -36.4853
26.3094 | -35.5866 | -36.4911
26.3155 | -35.587 | -36.4958
26.3216 | -35.5873 | -36.5016
26.3277 | -35.5877 | -36.5074
26.3338 | -35.5881 | -36.5131
26.3399 | -35.5885 | -36.5187
26.346 | -35.5888 | -36.524
26.3521 | -35.5892 | -36.5299
26.3582 | -35.5896 | -36.5357
26.3643 | -35.5899 | -36.5416
26.3704 | -35.5904 | -36.5472
26.3765 | -35.5907 | -36.5527
26.3826 | -35.5911 | -36.558
26.3887 | -35.5915 | -36.5636
26.3948 | -35.5919 | -36.569
26.4009 | -35.5923 | -36.575
26.407 | -35.5926 | -36.5806
26.4131 | -35.593 | -36.5865
26.4192 | -35.5933 | -36.5918
26.4253 | -35.5937 | -36.5973
26.4314 | -35.594 | -36.6028
26.4376 | -35.5944 | -36.6082
26.4437 | -35.5948 | -36.6136
26.4498 | -35.5952 | -36.6191
26.4559 | -35.5955 | -36.6244
26.462 | -35.5959 | -36.6304
26.4681 | -35.5963 | -36.6356
26.4742 | -35.5967 | -36.6413
26.4803 | -35.597 | -36.6466
26.4864 | -35.5973 | -36.6525
26.4925 | -35.5977 | -36.658
26.4986 | -35.5981 | -36.6631
26.5047 | -35.5985 | -36.6687
26.5108 | -35.5988 | -36.6743
26.5169 | -35.5991 | -36.6799
26.523 | -35.5993 | -36.6858
26.5291 | -35.5998 | -36.6913
26.5352 | -35.6003 | -36.697
26.5413 | -35.6007 | -36.7027
26.5474 | -35.6011 | -36.7082
26.5535 | -35.6014 | -36.714
26.5596 | -35.6017 | -36.7201
26.5657 | -35.6021 | -36.7256
26.5718 | -35.6026 | -36.7318
26.578 | -35.6029 | -36.7366
26.5841 | -35.6033 | -36.7422
26.5902 | -35.6036 | -36.7481
26.5963 | -35.604 | -36.7536
26.6024 | -35.6043 | -36.7594
26.6085 | -35.6047 | -36.7642
26.6146 | -35.6051 | -36.7699
26.6207 | -35.6056 | -36.7758
26.6268 | -35.6059 | -36.7816
26.6329 | -35.6062 | -36.7874
26.639 | -35.6066 | -36.7927
26.6451 | -35.607 | -36.798
26.6512 | -35.6073 | -36.8041
26.6573 | -35.6076 | -36.8103
26.6634 | -35.6079 | -36.8163
26.6695 | -35.6083 | -36.822
26.6756 | -35.6086 | -36.8278
26.6817 | -35.609 | -36.833
26.6878 | -35.6093 | -36.8382
26.6939 | -35.6096 | -36.843
26.7 | -35.6099 | -36.8479
26.7061 | -35.6102 | -36.8535
26.7122 | -35.6105 | -36.8598
26.7183 | -35.6108 | -36.8656
26.7245 | -35.6112 | -36.8708
26.7306 | -35.6116 | -36.8766
26.7367 | -35.6119 | -36.8819
26.7428 | -35.6123 | -36.8874
26.7489 | -35.6126 | -36.8934
26.755 | -35.613 | -36.8987
26.7611 | -35.6134 | -36.9042
26.7672 | -35.6138 | -36.9097
26.7733 | -35.6142 | -36.9152
26.7794 | -35.6146 | -36.9205
26.7855 | -35.615 | -36.9261
26.7916 | -35.6153 | -36.9317
26.7977 | -35.6157 | -36.937
26.8038 | -35.616 | -36.9428
26.8099 | -35.6164 | -36.9489
26.816 | -35.6168 | -36.9544
26.8221 | -35.6171 | -36.9596
26.8282 | -35.6175 | -36.9653
26.8343 | -35.6179 | -36.9705
26.8404 | -35.6183 | -36.976
26.8465 | -35.6186 | -36.9818
26.8526 | -35.619 | -36.9868
26.8587 | -35.6195 | -36.9926
26.8649 | -35.6199 | -36.9981
26.871 | -35.6202 | -37.0034
26.8771 | -35.6206 | -37.0092
26.8832 | -35.621 | -37.0144
26.8893 | -35.6214 | -37.0195
26.8954 | -35.6218 | -37.0253
26.9015 | -35.6222 | -37.0312
26.9076 | -35.6225 | -37.0368
26.9137 | -35.6229 | -37.0423
26.9198 | -35.6233 | -37.0479
26.9259 | -35.6236 | -37.0537
26.932 | -35.624 | -37.0593
26.9381 | -35.6243 | -37.0652
26.9442 | -35.6246 | -37.0709
26.9503 | -35.6249 | -37.0758
26.9564 | -35.6253 | -37.0808
26.9625 | -35.6256 | -37.0871
26.9686 | -35.626 | -37.0923
26.9747 | -35.6264 | -37.098
26.9808 | -35.6267 | -37.1035
26.9869 | -35.627 | -37.1087
26.993 | -35.6273 | -37.1142
26.9991 | -35.6277 | -37.12
27.0052 | -35.628 | -37.1255
27.0114 | -35.6284 | -37.131
27.0175 | -35.6288 | -37.1372
27.0236 | -35.6292 | -37.1421
27.0297 | -35.6296 | -37.1475
27.0358 | -35.63 | -37.1537
27.0419 | -35.6303 | -37.1596
27.048 | -35.6307 | -37.1651
27.0541 | -35.6311 | -37.1704
27.0602 | -35.6316 | -37.1762
27.0663 | -35.6319 | -37.182
27.0724 | -35.6322 | -37.1878
27.0785 | -35.6326 | -37.1936
27.0846 | -35.6329 | -37.1994
27.0907 | -35.6333 | -37.2047
27.0968 | -35.6336 | -37.2103
27.1029 | -35.634 | -37.2163
27.109 | -35.6343 | -37.2219
27.1151 | -35.6347 | -37.2272
27.1212 | -35.6351 | -37.2334
27.1273 | -35.6355 | -37.2396
27.1334 | -35.6359 | -37.2455
27.1395 | -35.6364 | -37.251
27.1456 | -35.6367 | -37.2564
27.1518 | -35.6371 | -37.2621
27.1579 | -35.6374 | -37.2669
27.164 | -35.6378 | -37.2725
27.1701 | -35.6382 | -37.2783
27.1762 | -35.6385 | -37.2838
27.1823 | -35.6389 | -37.2902
27.1884 | -35.6392 | -37.2966
27.1945 | -35.6396 | -37.3022
27.2006 | -35.6399 | -37.3074
27.2067 | -35.6403 | -37.3129
27.2128 | -35.6407 | -37.318
27.2189 | -35.641 | -37.3232
27.225 | -35.6414 | -37.3289
27.2311 | -35.6419 | -37.3351
27.2372 | -35.6422 | -37.3407
27.2433 | -35.6425 | -37.3466
27.2494 | -35.6429 | -37.3523
27.2555 | -35.6432 | -37.3577
27.2616 | -35.6435 | -37.3634
27.2677 | -35.6439 | -37.3689
27.2738 | -35.6442 | -37.3742
27.2799 | -35.6446 | -37.38
27.286 | -35.6448 | -37.3857
27.2921 | -35.6453 | -37.392
27.2983 | -35.6456 | -37.3978
27.3044 | -35.6459 | -37.404
27.3105 | -35.6463 | -37.4095
27.3166 | -35.6466 | -37.4148
27.3227 | -35.647 | -37.4201
27.3288 | -35.6473 | -37.4259
27.3349 | -35.6477 | -37.4309
27.341 | -35.648 | -37.4364
27.3471 | -35.6484 | -37.4421
27.3532 | -35.6488 | -37.4483
27.3593 | -35.6492 | -37.4541
27.3654 | -35.6495 | -37.4595
27.3715 | -35.6499 | -37.4649
27.3776 | -35.6503 | -37.4702
27.3837 | -35.6506 | -37.4758
27.3898 | -35.6511 | -37.4808
27.3959 | -35.6514 | -37.4866
27.402 | -35.6516 | -37.4923
27.4081 | -35.652 | -37.4974
27.4142 | -35.6522 | -37.5032
27.4203 | -35.6526 | -37.5089
27.4264 | -35.6529 | -37.5144
27.4325 | -35.6532 | -37.5199
27.4387 | -35.6535 | -37.5252
27.4448 | -35.6537 | -37.5307
27.4509 | -35.654 | -37.5368
27.457 | -35.6544 | -37.5418
27.4631 | -35.6548 | -37.5476
27.4692 | -35.6552 | -37.5523
27.4753 | -35.6556 | -37.5575
27.4814 | -35.6559 | -37.5634
27.4875 | -35.6563 | -37.569
27.4936 | -35.6567 | -37.575
27.4997 | -35.6572 | -37.5807
27.5058 | -35.6576 | -37.5864
27.5119 | -35.6578 | -37.5923
27.518 | -35.6581 | -37.5981
27.5241 | -35.6583 | -37.6037
27.5302 | -35.6586 | -37.6094
27.5363 | -35.659 | -37.6149
27.5424 | -35.6593 | -37.6213
27.5485 | -35.6596 | -37.627
27.5546 | -35.66 | -37.6324
27.5607 | -35.6603 | -37.6381
27.5668 | -35.6607 | -37.6443
27.5729 | -35.6611 | -37.6494
27.5791 | -35.6613 | -37.6553
27.5852 | -35.6618 | -37.6615
27.5913 | -35.6622 | -37.6666
27.5974 | -35.6625 | -37.6726
27.6035 | -35.6629 | -37.6787
27.6096 | -35.6631 | -37.6848
27.6157 | -35.6634 | -37.6903
27.6218 | -35.6638 | -37.696
27.6279 | -35.6641 | -37.7014
27.634 | -35.6645 | -37.7071
27.6401 | -35.6648 | -37.7121
27.6462 | -35.665 | -37.7183
27.6523 | -35.6655 | -37.7243
27.6584 | -35.6658 | -37.7305
27.6645 | -35.6662 | -37.7367
27.6706 | -35.6666 | -37.7424
27.6767 | -35.6669 | -37.7489
27.6828 | -35.6673 | -37.7544
27.6889 | -35.6676 | -37.7602
27.695 | -35.6679 | -37.7653
27.7011 | -35.6682 | -37.7706
27.7072 | -35.6686 | -37.7757
27.7133 | -35.6689 | -37.7818
27.7194 | -35.6692 | -37.7869
27.7256 | -35.6696 | -37.7926
27.7317 | -35.6699 | -37.7977
27.7378 | -35.6702 | -37.8037
27.7439 | -35.6705 | -37.8091
27.75 | -35.6708 | -37.8144
27.7561 | -35.671 | -37.8187
27.7622 | -35.6713 | -37.8248
27.7683 | -35.6716 | -37.8308
27.7744 | -35.6719 | -37.8362
27.7805 | -35.6722 | -37.842
27.7866 | -35.6726 | -37.8472
27.7927 | -35.673 | -37.8516
27.7988 | -35.6733 | -37.8573
27.8049 | -35.6736 | -37.8631
27.811 | -35.674 | -37.8687
27.8171 | -35.6745 | -37.8745
27.8232 | -35.6748 | -37.8805
27.8293 | -35.6751 | -37.8861
27.8354 | -35.6755 | -37.8919
27.8415 | -35.676 | -37.8975
27.8476 | -35.6765 | -37.9031
27.8537 | -35.6768 | -37.9087
27.8598 | -35.6772 | -37.9126
27.866 | -35.6775 | -37.918
27.8721 | -35.6779 | -37.9234
27.8782 | -35.6783 | -37.9291
27.8843 | -35.6787 | -37.9344
27.8904 | -35.6791 | -37.9402
27.8965 | -35.6795 | -37.9453
27.9026 | -35.6798 | -37.9512
27.9087 | -35.6803 | -37.9568
27.9148 | -35.6806 | -37.9622
27.9209 | -35.6809 | -37.9673
27.927 | -35.6813 | -37.973
27.9331 | -35.6816 | -37.9776
27.9392 | -35.682 | -37.9835
27.9453 | -35.6823 | -37.9887
27.9514 | -35.6825 | -37.9939
27.9575 | -35.6829 | -37.9996
27.9636 | -35.6832 | -38.0054
27.9697 | -35.6836 | -38.0108
27.9758 | -35.684 | -38.0168
27.9819 | -35.6843 | -38.023
27.988 | -35.6847 | -38.0287
27.9941 | -35.6851 | -38.0339
28.0002 | -35.6853 | -38.0401
28.0063 | -35.6857 | -38.0455
28.0125 | -35.6861 | -38.0508
28.0186 | -35.6866 | -38.0564
28.0247 | -35.687 | -38.0607
28.0308 | -35.6873 | -38.0659
28.0369 | -35.6876 | -38.0715
28.043 | -35.6879 | -38.0775
28.0491 | -35.6883 | -38.0837
28.0552 | -35.6887 | -38.0888
28.0613 | -35.6891 | -38.0939
28.0674 | -35.6894 | -38.1003
28.0735 | -35.6897 | -38.1077

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.48 | -43.2127
21.4908 | -35.5098 | -35.2297
16.4948 | -34.8913 | -27.7899
12.6602 | -34.5279 | -21.4789
9.7171 | -34.3132 | -16.557
7.4582 | -34.1632 | -12.5638
5.7244 | -34.0575 | -9.3244
4.3936 | -33.9914 | -6.7158
3.3722 | -33.9739 | -4.6619
2.5883 | -33.9751 | -3.1017
1.9866 | -33.9791 | -1.853
1.5248 | -33.9749 | -0.85261
1.1703 | -33.9595 | -0.19947
1 | -33.9648 | 0.12481
0.89822 | -33.953 | 0.3997
0.68941 | -33.9448 | 0.98427
0.52919 | -33.9462 | 1.4432
0.40612 | -33.9513 | 2.0518
0.31171 | -33.9431 | 2.5099
0.23928 | -33.9256 | 3.2514
0.18358 | -33.9492 | 3.669
0.14091 | -33.9689 | 4.5531
0.10819 | -33.9721 | 5.7303
0.083017 | -33.9966 | 7.2173
0.063705 | -34.0526 | 9.1031
0.048923 | -34.1165 | 11.6202
0.037479 | -34.2246 | 14.8333
0.028801 | -34.4161 | 18.8688
0.022078 | -34.7144 | 24.2125
0.017023 | -35.1951 | 30.6304
0.013018 | -36.0067 | 38.3279
0.010014 | -37.1191 | 46.2175

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003751 | -5.3936
0.007501 | -26.2175
0.011252 | -35.006
0.015002 | -40.073
0.018753 | -43.5381
0.022504 | -45.7884
0.026254 | -46.9333
0.030005 | -47.9215
0.033755 | -49.6146
0.037506 | -53.09
0.041256 | -54.0809
0.045007 | -53.9061
0.048758 | -55.2622
0.052508 | -56.3571
0.056259 | -56.4001
0.060009 | -56.8302
0.06376 | -57.778
0.067511 | -58.511
0.071261 | -59.3814
0.075012 | -59.7549
0.078762 | -60.3945
0.082513 | -62.3819
0.086263 | -63.471
0.090014 | -62.3778
0.093765 | -61.7275
0.097515 | -61.8789
0.10127 | -62.9117
0.10502 | -62.9407
0.10877 | -63.0528
0.11252 | -62.9469
0.11627 | -63.395
0.12002 | -64.4567
0.12377 | -66.0838
0.12752 | -66.3566
0.13127 | -66.0431
0.13502 | -65.8678
0.13877 | -66.5714
0.14252 | -66.8563
0.14627 | -67.2776
0.15002 | -67.6383
0.15377 | -67.914
0.15752 | -68.7158
0.16128 | -68.9372
0.16503 | -69.4103
0.16878 | -69.6134
0.17253 | -69.3793
0.17628 | -69.2129
0.18003 | -69.511
0.18378 | -69.1597
0.18753 | -69.2407
0.19128 | -69.3495
0.19503 | -69.7628
0.19878 | -70.5132
0.20253 | -70.1939
0.20628 | -70.0396
0.21003 | -70.3395
0.21378 | -70.5078
0.21753 | -70.1715
0.22129 | -70.6281
0.22504 | -71.8443
0.22879 | -72.0591
0.23254 | -72.4384
0.23629 | -72.8579
0.24004 | -72.6641
0.24379 | -72.7802
0.24754 | -73.6122
0.25129 | -73.394
0.25504 | -73.0312
0.25879 | -72.8706
0.26254 | -72.139
0.26629 | -72.1493
0.27004 | -73.3896
0.27379 | -74.2971
0.27754 | -74.5065
0.28129 | -75.6008
0.28504 | -74.1947
0.2888 | -74.4008
0.29255 | -75.1411
0.2963 | -74.6147
0.30005 | -75.0456
0.3038 | -75.5281
0.30755 | -75.5578
0.3113 | -74.8268
0.31505 | -75.0142
0.3188 | -74.477
0.32255 | -74.3704
0.3263 | -74.746
0.33005 | -76.6178
0.3338 | -76.5084
0.33755 | -74.3813
0.3413 | -74.5596
0.34505 | -75.4646
0.3488 | -75.8073
0.35256 | -75.3902
0.35631 | -76.0249
0.36006 | -75.8533
0.36381 | -75.6102
0.36756 | -75.9579
0.37131 | -75.5049
0.37506 | -76.0905
0.37881 | -77.2749
0.38256 | -76.6563
0.38631 | -76.7777
0.39006 | -77.6603
0.39381 | -77.1953
0.39756 | -76.7679
0.40131 | -76.6395
0.40506 | -76.3528
0.40881 | -76.6186
0.41256 | -76.3667
0.41631 | -76.5416
0.42007 | -76.8176
0.42382 | -76.5224
0.42757 | -77.0633
0.43132 | -79.5888
0.43507 | -79.1967
0.43882 | -77.9173
0.44257 | -77.5916
0.44632 | -77.6394
0.45007 | -78.1823
0.45382 | -78.9167
0.45757 | -78.0497
0.46132 | -77.7377
0.46507 | -78.0069
0.46882 | -78.7031
0.47257 | -78.8081
0.47632 | -78.416
0.48007 | -78.9648
0.48383 | -78.5512
0.48758 | -77.7824
0.49133 | -78.5718
0.49508 | -79.5242
0.49883 | -79.106
0.50258 | -78.5983
0.50633 | -78.6482
0.51008 | -79.5004
0.51383 | -79.4795
0.51758 | -79.7109
0.52133 | -80.5835
0.52508 | -79.9154
0.52883 | -78.8535
0.53258 | -79.0882
0.53633 | -79.3911
0.54008 | -79.339
0.54383 | -79.086
0.54759 | -78.7983
0.55134 | -79.0244
0.55509 | -79.9741
0.55884 | -80.7718
0.56259 | -80.4925
0.56634 | -80.2059
0.57009 | -80.3135
0.57384 | -80.5228
0.57759 | -80.5713
0.58134 | -80.5361
0.58509 | -80.6053
0.58884 | -80.8436
0.59259 | -80.8887
0.59634 | -80.7455
0.60009 | -80.9042
0.60384 | -81.3675
0.60759 | -81.7288
0.61135 | -81.7452
0.6151 | -81.3472
0.61885 | -80.8977
0.6226 | -80.6404
0.62635 | -80.5327
0.6301 | -80.5686
0.63385 | -80.7786
0.6376 | -80.9188
0.64135 | -81.0332
0.6451 | -81.3179
0.64885 | -81.511
0.6526 | -81.323
0.65635 | -80.9224
0.6601 | -80.6579
0.66385 | -80.5949
0.6676 | -80.5736
0.67136 | -80.7332
0.6751 | -81.1245
0.67886 | -81.7651
0.68261 | -82.0748
0.68636 | -81.4167
0.69011 | -81.028
0.69386 | -81.3932
0.69761 | -81.8491
0.70136 | -82.0991
0.70511 | -82.3399
0.70886 | -82.5975
0.71261 | -82.5986
0.71636 | -82.2769
0.72011 | -81.9267
0.72386 | -81.8482
0.72761 | -82.0735
0.73136 | -82.1607
0.73511 | -82.0347
0.73886 | -82.14
0.74262 | -82.6712
0.74637 | -82.9572
0.75012 | -82.9018
0.75387 | -82.7797
0.75762 | -82.6305
0.76137 | -82.947
0.76512 | -83.6386
0.76887 | -84.0595
0.77262 | -83.9423
0.77637 | -83.6116
0.78012 | -83.3142
0.78387 | -83.4422
0.78762 | -83.792
0.79137 | -83.4012
0.79512 | -82.6981
0.79888 | -82.6754
0.80263 | -82.8765
0.80638 | -82.5564
0.81013 | -82.5428
0.81388 | -82.9611
0.81763 | -83.0057
0.82138 | -82.7126
0.82513 | -82.6512
0.82888 | -82.7038
0.83263 | -82.8237
0.83638 | -82.9785
0.84013 | -83.0421
0.84388 | -83.2733
0.84763 | -83.7211
0.85138 | -84.0161
0.85513 | -84.0339
0.85888 | -83.9127
0.86264 | -84.1519
0.86638 | -84.7119
0.87014 | -85.0221
0.87389 | -85.1068
0.87764 | -84.9293
0.88139 | -84.3115
0.88514 | -84.0497
0.88889 | -83.9292
0.89264 | -83.7342
0.89639 | -83.8251
0.90014 | -84.3075
0.90389 | -84.4448
0.90764 | -83.9914
0.91139 | -83.6779
0.91514 | -83.77
0.91889 | -83.8151
0.92264 | -83.6202
0.92639 | -83.7138
0.93014 | -84.1097
0.9339 | -84.4847
0.93765 | -84.6751
0.9414 | -84.7757
0.94515 | -84.9244
0.9489 | -84.8485
0.95265 | -84.7553
0.9564 | -84.7399
0.96015 | -84.8131
0.9639 | -85.0721
0.96765 | -85.2775
0.9714 | -84.8882
0.97515 | -84.3682
0.9789 | -84.4671
0.98265 | -84.8359
0.9864 | -84.9662
0.99016 | -84.9354
0.99391 | -84.8441
0.99766 | -84.479
1.0014 | -84.4251
1.0052 | -84.9753
1.0089 | -85.7521
1.0127 | -85.9169
1.0164 | -85.6856
1.0202 | -85.4865
1.0239 | -85.4046
1.0277 | -85.3557
1.0314 | -85.2076
1.0352 | -85.3198
1.0389 | -85.9226
1.0427 | -86.1639
1.0464 | -85.747
1.0502 | -85.2264
1.0539 | -84.9435
1.0577 | -85.0514
1.0614 | -85.4008
1.0652 | -85.5887
1.0689 | -85.6203
1.0727 | -85.7615
1.0764 | -86.0232
1.0802 | -86.1996
1.0839 | -86.11
1.0877 | -85.8983
1.0914 | -85.9325
1.0952 | -86.2493
1.0989 | -86.4882
1.1027 | -86.3589
1.1064 | -86.032
1.1102 | -85.8553
1.1139 | -85.9768
1.1177 | -86.2655
1.1214 | -86.4708
1.1252 | -86.581
1.1289 | -86.8027
1.1327 | -87.0966
1.1364 | -87.1753
1.1402 | -86.9346
1.1439 | -86.5503
1.1477 | -86.2203
1.1514 | -86.0977
1.1552 | -86.2988
1.1589 | -86.7504
1.1627 | -87.1645
1.1664 | -87.3008
1.1702 | -87.2617
1.1739 | -87.3373
1.1777 | -87.5615
1.1814 | -87.7603
1.1852 | -87.777
1.1889 | -87.4798
1.1927 | -87.0555
1.1964 | -86.9064
1.2002 | -87.1408
1.2039 | -87.4701
1.2077 | -87.5449
1.2114 | -87.3471
1.2152 | -87.1196
1.2189 | -86.937
1.2227 | -86.6518
1.2264 | -86.3215
1.2302 | -86.1965
1.2339 | -86.3438
1.2377 | -86.5987
1.2414 | -86.7461
1.2452 | -86.7146
1.2489 | -86.6148
1.2527 | -86.6566
1.2564 | -86.9923
1.2602 | -87.502
1.2639 | -87.867
1.2677 | -87.9563
1.2714 | -87.9543
1.2752 | -88.033
1.279 | -88.116
1.2827 | -87.9936
1.2865 | -87.613
1.2902 | -87.1576
1.294 | -86.8604
1.2977 | -86.822
1.3015 | -86.969
1.3052 | -87.0387
1.309 | -86.8414
1.3127 | -86.5959
1.3165 | -86.6092
1.3202 | -86.926
1.324 | -87.3799
1.3277 | -87.6946
1.3315 | -87.6277
1.3352 | -87.2976
1.339 | -87.0528
1.3427 | -86.9759
1.3465 | -86.8919
1.3502 | -86.7092
1.354 | -86.5467
1.3577 | -86.5572
1.3615 | -86.8025
1.3652 | -87.2554
1.369 | -87.7978
1.3727 | -88.2547
1.3765 | -88.5586
1.3802 | -88.6617
1.384 | -88.4582
1.3877 | -88.0952
1.3915 | -87.8736
1.3952 | -87.8515
1.399 | -87.9001
1.4027 | -87.9207
1.4065 | -87.9357
1.4102 | -87.998
1.414 | -88.0711
1.4177 | -88.1208
1.4215 | -88.1805
1.4252 | -88.2501
1.429 | -88.2957
1.4327 | -88.2988
1.4365 | -88.2273
1.4402 | -88.1055
1.444 | -88.059
1.4477 | -88.0809
1.4515 | -88.0416
1.4552 | -88.0193
1.459 | -88.1088
1.4627 | -88.1971
1.4665 | -88.2943
1.4702 | -88.5711
1.474 | -88.9915
1.4777 | -89.2715
1.4815 | -89.1795
1.4852 | -88.8527
1.489 | -88.6892
1.4927 | -88.8256
1.4965 | -88.9304
1.5002 | -88.7177
1.504 | -88.4361
1.5077 | -88.3748
1.5115 | -88.5539
1.5152 | -88.8644
1.519 | -89.1377
1.5227 | -89.2618
1.5265 | -89.2969
1.5302 | -89.3329
1.534 | -89.3814
1.5377 | -89.356
1.5415 | -89.1673
1.5452 | -88.9116
1.549 | -88.7521
1.5527 | -88.777
1.5565 | -88.9921
1.5602 | -89.2368
1.564 | -89.2786
1.5677 | -89.0669
1.5715 | -88.7461
1.5752 | -88.5504
1.579 | -88.6281
1.5827 | -88.9102
1.5865 | -89.173
1.5902 | -89.3362
1.594 | -89.5309
1.5978 | -89.7301
1.6015 | -89.7191
1.6053 | -89.4599
1.609 | -89.1926
1.6128 | -89.0797
1.6165 | -89.0093
1.6203 | -88.8405
1.624 | -88.7761
1.6278 | -88.8564
1.6315 | -89.0535
1.6353 | -89.2519
1.639 | -89.3302
1.6428 | -89.285
1.6465 | -89.1974
1.6503 | -89.1352
1.654 | -89.1035
1.6578 | -89.0943
1.6615 | -89.1659
1.6653 | -89.3692
1.669 | -89.6348
1.6728 | -89.8256
1.6765 | -89.8614
1.6803 | -89.7884
1.684 | -89.7119
1.6878 | -89.6983
1.6915 | -89.7304
1.6953 | -89.7533
1.699 | -89.7206
1.7028 | -89.6059
1.7065 | -89.4404
1.7103 | -89.3236
1.714 | -89.3573
1.7178 | -89.5696
1.7215 | -89.8905
1.7253 | -90.1765
1.729 | -90.2909
1.7328 | -90.21
1.7365 | -89.9737
1.7403 | -89.6567
1.744 | -89.4045
1.7478 | -89.3572
1.7515 | -89.5368
1.7553 | -89.858
1.759 | -90.2288
1.7628 | -90.5773
1.7665 | -90.7965
1.7703 | -90.7896
1.774 | -90.574
1.7778 | -90.2886
1.7815 | -90.0796
1.7853 | -90.0049
1.789 | -90.0517
1.7928 | -90.1727
1.7965 | -90.3186
1.8003 | -90.4367
1.804 | -90.4744
1.8078 | -90.4127
1.8115 | -90.3359
1.8153 | -90.35
1.819 | -90.4778
1.8228 | -90.6218
1.8265 | -90.6958
1.8303 | -90.7081
1.834 | -90.6913
1.8378 | -90.5906
1.8415 | -90.3809
1.8453 | -90.1627
1.849 | -90.0837
1.8528 | -90.1682
1.8565 | -90.341
1.8603 | -90.5083
1.864 | -90.5937
1.8678 | -90.5613
1.8715 | -90.4868
1.8753 | -90.4574
1.879 | -90.5088
1.8828 | -90.5987
1.8865 | -90.62
1.8903 | -90.4547
1.894 | -90.1313
1.8978 | -89.8621
1.9015 | -89.8192
1.9053 | -89.9876
1.909 | -90.2254
1.9128 | -90.3752
1.9165 | -90.4066
1.9203 | -90.407
1.9241 | -90.4379
1.9278 | -90.444
1.9316 | -90.3371
1.9353 | -90.1823
1.9391 | -90.1562
1.9428 | -90.3372
1.9466 | -90.6384
1.9503 | -90.8981
1.9541 | -91.0054
1.9578 | -90.9623
1.9616 | -90.8923
1.9653 | -90.9303
1.9691 | -91.1279
1.9728 | -91.4053
1.9766 | -91.6511
1.9803 | -91.8282
1.9841 | -91.9402
1.9878 | -91.9604
1.9916 | -91.859
1.9953 | -91.6834
1.9991 | -91.5386
2.0028 | -91.4241
2.0066 | -91.2854
2.0103 | -91.1142
2.0141 | -90.9816
2.0178 | -90.9481
2.0216 | -90.9972
2.0253 | -91.0743
2.0291 | -91.1194
2.0328 | -91.0927
2.0366 | -91.0112
2.0403 | -90.938
2.0441 | -90.9288
2.0478 | -90.9902
2.0516 | -91.1031
2.0553 | -91.2295
2.0591 | -91.2994
2.0628 | -91.2507
2.0666 | -91.126
2.0703 | -91.0609
2.0741 | -91.1196
2.0778 | -91.2361
2.0816 | -91.2994
2.0853 | -91.277
2.0891 | -91.2306
2.0928 | -91.2097
2.0966 | -91.264
2.1003 | -91.4321
2.1041 | -91.6853
2.1078 | -91.9025
2.1116 | -91.9313
2.1153 | -91.7442
2.1191 | -91.4833
2.1228 | -91.3478
2.1266 | -91.4632
2.1303 | -91.7754
2.1341 | -92.0932
2.1378 | -92.267
2.1416 | -92.327
2.1453 | -92.3711
2.1491 | -92.3798
2.1528 | -92.2689
2.1566 | -92.0407
2.1603 | -91.763
2.1641 | -91.5235
2.1678 | -91.3229
2.1716 | -91.1715
2.1753 | -91.0938
2.1791 | -91.0936
2.1828 | -91.1503
2.1866 | -91.248
2.1903 | -91.3549
2.1941 | -91.4553
2.1978 | -91.5545
2.2016 | -91.6504
2.2053 | -91.7572
2.2091 | -91.8909
2.2128 | -92.0393
2.2166 | -92.1561
2.2203 | -92.208
2.2241 | -92.1718
2.2278 | -92.0552
2.2316 | -91.9104
2.2353 | -91.7987
2.2391 | -91.7647
2.2428 | -91.8115
2.2466 | -91.8976
2.2504 | -91.9629
2.2541 | -91.9733
2.2579 | -91.9414
2.2616 | -91.903
2.2654 | -91.8972
2.2691 | -91.9233
2.2729 | -91.9592
2.2766 | -91.9668
2.2804 | -91.9437
2.2841 | -91.9189
2.2879 | -91.9331
2.2916 | -92.0284
2.2954 | -92.183
2.2991 | -92.3185
2.3029 | -92.3711
2.3066 | -92.3658
2.3104 | -92.3794
2.3141 | -92.4706
2.3179 | -92.6251
2.3216 | -92.768
2.3254 | -92.8438
2.3291 | -92.8464
2.3329 | -92.8027
2.3366 | -92.7488
2.3404 | -92.6929
2.3441 | -92.628
2.3479 | -92.5441
2.3516 | -92.4475
2.3554 | -92.3462
2.3591 | -92.2765
2.3629 | -92.2579
2.3666 | -92.3054
2.3704 | -92.434
2.3741 | -92.6209
2.3779 | -92.823
2.3816 | -92.9744
2.3854 | -93.0191
2.3891 | -92.9495
2.3929 | -92.8005
2.3966 | -92.6222
2.4004 | -92.4671
2.4041 | -92.3875
2.4079 | -92.4001
2.4116 | -92.4909
2.4154 | -92.6307
2.4191 | -92.7766
2.4229 | -92.8844
2.4266 | -92.9428
2.4304 | -92.9638
2.4341 | -92.974
2.4379 | -92.9676
2.4416 | -92.925
2.4454 | -92.8617
2.4491 | -92.8316
2.4529 | -92.8681
2.4566 | -92.9527
2.4604 | -93.0331
2.4641 | -93.082
2.4679 | -93.1156
2.4716 | -93.1667
2.4754 | -93.2281
2.4791 | -93.2746
2.4829 | -93.2801
2.4866 | -93.241
2.4904 | -93.1756
2.4941 | -93.1244
2.4979 | -93.1106
2.5016 | -93.147
2.5054 | -93.2332
2.5091 | -93.3256
2.5129 | -93.3815
2.5166 | -93.3911
2.5204 | -93.3546
2.5241 | -93.2918
2.5279 | -93.223
2.5316 | -93.1662
2.5354 | -93.1336
2.5391 | -93.1203
2.5429 | -93.0904
2.5466 | -93.0059
2.5504 | -92.8706
2.5541 | -92.7352
2.5579 | -92.6736
2.5617 | -92.7282
2.5654 | -92.8683
2.5692 | -93.0248
2.5729 | -93.1389
2.5767 | -93.1895
2.5804 | -93.1884
2.5842 | -93.1431
2.5879 | -93.0634
2.5917 | -92.9721
2.5954 | -92.9145
2.5992 | -92.9426
2.6029 | -93.0889
2.6067 | -93.3476
2.6104 | -93.6671
2.6142 | -93.9557
2.6179 | -94.1159
2.6217 | -94.0814
2.6254 | -93.8482
2.6292 | -93.4649
2.6329 | -93.0415
2.6367 | -92.705
2.6404 | -92.5444
2.6442 | -92.5683
2.6479 | -92.7049
2.6517 | -92.8485
2.6554 | -92.9302
2.6592 | -92.9625
2.6629 | -93.0012
2.6667 | -93.0725
2.6704 | -93.1629
2.6742 | -93.2447
2.6779 | -93.3013
2.6817 | -93.3447
2.6854 | -93.3908
2.6892 | -93.4487
2.6929 | -93.5232
2.6967 | -93.6141
2.7004 | -93.7287
2.7042 | -93.8927
2.7079 | -94.0511
2.7117 | -94.2128
2.7154 | -94.3568
2.7192 | -94.453
2.7229 | -94.4936
2.7267 | -94.4628
2.7304 | -94.365
2.7342 | -94.2102
2.7379 | -94.0262
2.7417 | -93.8499
2.7454 | -93.705
2.7492 | -93.6028
2.7529 | -93.5214
2.7567 | -93.4372
2.7604 | -93.3367
2.7642 | -93.2227
2.7679 | -93.1198
2.7717 | -93.0408
2.7754 | -92.9984
2.7792 | -93.0045
2.7829 | -93.0706
2.7867 | -93.1981
2.7904 | -93.3658
2.7942 | -93.5283
2.7979 | -93.6321
2.8017 | -93.6562
2.8054 | -93.6102
2.8092 | -93.5256
2.8129 | -93.4505
2.8167 | -93.3998
2.8204 | -93.3789
2.8242 | -93.3869
2.8279 | -93.4196
2.8317 | -93.4893
2.8354 | -93.6066
2.8392 | -93.7642
2.8429 | -93.933
2.8467 | -94.0772
2.8504 | -94.1748
2.8542 | -94.2158
2.8579 | -94.2151
2.8617 | -94.1654
2.8654 | -94.086
2.8692 | -93.9997
2.8729 | -93.9251
2.8767 | -93.8787
2.8805 | -93.8784
2.8842 | -93.9007
2.888 | -93.9195
2.8917 | -93.9077
2.8955 | -93.8576
2.8992 | -93.7769
2.903 | -93.6889
2.9067 | -93.6161
2.9105 | -93.5748
2.9142 | -93.5716
2.918 | -93.5975
2.9217 | -93.6427
2.9255 | -93.695
2.9292 | -93.744
2.933 | -93.8146
2.9367 | -93.9116
2.9405 | -94.0191
2.9442 | -94.1141
2.948 | -94.1744
2.9517 | -94.1945
2.9555 | -94.1743
2.9592 | -94.1235
2.963 | -94.0308
2.9667 | -93.8873
2.9705 | -93.7029
2.9742 | -93.5191
2.978 | -93.3869
2.9817 | -93.3381
2.9855 | -93.3941
2.9892 | -93.5506
2.993 | -93.776
2.9967 | -94.0287
3.0005 | -94.262
3.0042 | -94.4392
3.008 | -94.5407
3.0117 | -94.576
3.0155 | -94.5785
3.0192 | -94.5704
3.023 | -94.5606
3.0267 | -94.5484
3.0305 | -94.5535
3.0342 | -94.5763
3.038 | -94.6211
3.0417 | -94.6846
3.0455 | -94.764
3.0492 | -94.85
3.053 | -94.9445
3.0567 | -95.055
3.0605 | -95.1732
3.0642 | -95.2686
3.068 | -95.3159
3.0717 | -95.3026
3.0755 | -95.2453
3.0792 | -95.1486
3.083 | -95.0417
3.0867 | -94.939
3.0905 | -94.8572
3.0942 | -94.8067
3.098 | -94.8077
3.1017 | -94.8515
3.1055 | -94.9124
3.1092 | -94.9684
3.113 | -94.9917
3.1167 | -94.9687
3.1205 | -94.9011
3.1242 | -94.8075
3.128 | -94.7232
3.1317 | -94.6663
3.1355 | -94.6568
3.1392 | -94.6879
3.143 | -94.7405
3.1467 | -94.789
3.1505 | -94.8037
3.1542 | -94.7786
3.158 | -94.7277
3.1617 | -94.6796
3.1655 | -94.6664
3.1692 | -94.7148
3.173 | -94.8129
3.1767 | -94.9295
3.1805 | -95.0319
3.1842 | -95.0896
3.188 | -95.0942
3.1917 | -95.0623
3.1955 | -95.0072
3.1993 | -94.9551
3.203 | -94.9085
3.2068 | -94.8654
3.2105 | -94.8435
3.2143 | -94.8485
3.218 | -94.8743
3.2218 | -94.9112
3.2255 | -94.9315
3.2293 | -94.9184
3.233 | -94.8806
3.2368 | -94.8401
3.2405 | -94.8125
3.2443 | -94.8508
3.248 | -94.8818
3.2518 | -94.9299
3.2555 | -94.991
3.2593 | -95.0643
3.263 | -95.1297
3.2668 | -95.1826
3.2705 | -95.2241
3.2743 | -95.2557
3.278 | -95.2881
3.2818 | -95.3281
3.2855 | -95.3812
3.2893 | -95.4537
3.293 | -95.5407
3.2968 | -95.6286
3.3005 | -95.706
3.3043 | -95.7571
3.308 | -95.758
3.3118 | -95.712
3.3155 | -95.629
3.3193 | -95.5367
3.323 | -95.4498
3.3268 | -95.3691
3.3305 | -95.3016
3.3343 | -95.2403
3.338 | -95.1769
3.3418 | -95.1106
3.3455 | -95.0564
3.3493 | -95.0257
3.353 | -95.0301
3.3568 | -95.0767
3.3605 | -95.1689
3.3643 | -95.3069
3.368 | -95.4714
3.3718 | -95.6452
3.3755 | -95.8064
3.3793 | -95.946
3.383 | -96.046
3.3868 | -96.0918
3.3905 | -96.1094
3.3943 | -96.1081
3.398 | -96.0872
3.4018 | -96.0449
3.4055 | -95.9978
3.4093 | -95.9469
3.413 | -95.8944
3.4168 | -95.8469
3.4205 | -95.8133
3.4243 | -95.7891
3.428 | -95.7769
3.4318 | -95.7744
3.4355 | -95.7776
3.4393 | -95.7953
3.443 | -95.8283
3.4468 | -95.8627
3.4505 | -95.8977
3.4543 | -95.9317
3.458 | -95.9555
3.4618 | -95.9567
3.4655 | -95.9479
3.4693 | -95.9406
3.473 | -95.9362
3.4768 | -95.9322
3.4805 | -95.9298
3.4843 | -95.909
3.488 | -95.8657
3.4918 | -95.812
3.4955 | -95.7664
3.4993 | -95.7476
3.503 | -95.7757
3.5068 | -95.8576
3.5105 | -95.9768
3.5143 | -96.1019
3.5181 | -96.1984
3.5218 | -96.2224
3.5256 | -96.1483
3.5293 | -96.0028
3.5331 | -95.8286
3.5368 | -95.6694
3.5406 | -95.5557
3.5443 | -95.4972
3.5481 | -95.4868
3.5518 | -95.5114
3.5556 | -95.5532
3.5593 | -95.6077
3.5631 | -95.6813
3.5668 | -95.7636
3.5706 | -95.8461
3.5743 | -95.9276
3.5781 | -95.9953
3.5818 | -96.0481
3.5856 | -96.081
3.5893 | -96.1036
3.5931 | -96.1152
3.5968 | -96.1211
3.6006 | -96.127
3.6043 | -96.1378
3.6081 | -96.1586
3.6118 | -96.1877
3.6156 | -96.2033
3.6193 | -96.1992
3.6231 | -96.1821
3.6268 | -96.159
3.6306 | -96.1313
3.6343 | -96.1198
3.6381 | -96.13
3.6418 | -96.1614
3.6456 | -96.2164
3.6493 | -96.2915
3.6531 | -96.3569
3.6568 | -96.3945
3.6606 | -96.4046
3.6643 | -96.3956
3.6681 | -96.3902
3.6718 | -96.4125
3.6756 | -96.4594
3.6793 | -96.5134
3.6831 | -96.5439
3.6868 | -96.5455
3.6906 | -96.5192
3.6943 | -96.4808
3.6981 | -96.4566
3.7018 | -96.455
3.7056 | -96.4837
3.7093 | -96.5434
3.7131 | -96.6106
3.7168 | -96.6614
3.7206 | -96.6824
3.7243 | -96.6719
3.7281 | -96.6378
3.7318 | -96.6019
3.7356 | -96.586
3.7393 | -96.6019
3.7431 | -96.6359
3.7468 | -96.6901
3.7506 | -96.7583
3.7543 | -96.8253
3.7581 | -96.8797
3.7618 | -96.9037
3.7656 | -96.9032
3.7693 | -96.8832
3.7731 | -96.8527
3.7768 | -96.8129
3.7806 | -96.765
3.7843 | -96.7139
3.7881 | -96.6771
3.7918 | -96.6518
3.7956 | -96.6432
3.7993 | -96.6584
3.8031 | -96.7053
3.8068 | -96.7696
3.8106 | -96.8456
3.8143 | -96.9244
3.8181 | -96.9912
3.8218 | -97.0318
3.8256 | -97.0255
3.8293 | -96.9744
3.8331 | -96.8855
3.8369 | -96.7755
3.8406 | -96.656
3.8444 | -96.5512
3.8481 | -96.4699
3.8519 | -96.4254
3.8556 | -96.4255
3.8594 | -96.4707
3.8631 | -96.554
3.8669 | -96.6675
3.8706 | -96.7983
3.8744 | -96.9199
3.8781 | -97.0103
3.8819 | -97.0627
3.8856 | -97.0755
3.8894 | -97.0569
3.8931 | -97.0221
3.8969 | -96.9896
3.9006 | -96.9679
3.9044 | -96.9611
3.9081 | -96.9662
3.9119 | -96.9728
3.9156 | -96.9701
3.9194 | -96.9436
3.9231 | -96.8787
3.9269 | -96.778
3.9306 | -96.6546
3.9344 | -96.5241
3.9381 | -96.4096
3.9419 | -96.3302
3.9456 | -96.2958
3.9494 | -96.3122
3.9531 | -96.3828
3.9569 | -96.4988
3.9606 | -96.6424
3.9644 | -96.8075
3.9681 | -96.9734
3.9719 | -97.1103
3.9756 | -97.1995
3.9794 | -97.2347
3.9831 | -97.2288
3.9869 | -97.1943
3.9906 | -97.1525
3.9944 | -97.1107
3.9981 | -97.0794
4.0019 | -97.0516
4.0056 | -97.0171
4.0094 | -96.9735
4.0131 | -96.9257
4.0169 | -96.8796
4.0206 | -96.852
4.0244 | -96.8525
4.0281 | -96.8848
4.0319 | -96.9546
4.0356 | -97.0667
4.0394 | -97.2045
4.0431 | -97.3538
4.0469 | -97.4944
4.0506 | -97.6107
4.0544 | -97.6787
4.0581 | -97.7007
4.0619 | -97.6784
4.0656 | -97.6193
4.0694 | -97.5388
4.0731 | -97.4513
4.0769 | -97.3763
4.0806 | -97.3226
4.0844 | -97.2928
4.0881 | -97.2819
4.0919 | -97.2882
4.0956 | -97.311
4.0994 | -97.3282
4.1031 | -97.3444
4.1069 | -97.3585
4.1106 | -97.3638
4.1144 | -97.3631
4.1181 | -97.3586
4.1219 | -97.3444
4.1256 | -97.3201
4.1294 | -97.3033
4.1331 | -97.3011
4.1369 | -97.3168
4.1406 | -97.3401
4.1444 | -97.3708
4.1481 | -97.3944
4.1519 | -97.4062
4.1556 | -97.3889
4.1594 | -97.3517
4.1631 | -97.3206
4.1669 | -97.2997
4.1707 | -97.2922
4.1744 | -97.3055
4.1782 | -97.3318
4.1819 | -97.3771
4.1857 | -97.4398
4.1894 | -97.5084
4.1932 | -97.5693
4.1969 | -97.6205
4.2007 | -97.6657
4.2044 | -97.7107
4.2082 | -97.756
4.2119 | -97.8119
4.2157 | -97.8789
4.2194 | -97.9512
4.2232 | -98.01
4.2269 | -98.0475
4.2307 | -98.0615
4.2344 | -98.0463
4.2382 | -98.0079
4.2419 | -97.9406
4.2457 | -97.8364
4.2494 | -97.7148
4.2532 | -97.6012
4.2569 | -97.5032
4.2607 | -97.4298
4.2644 | -97.3853
4.2682 | -97.3748
4.2719 | -97.3975
4.2757 | -97.4442
4.2794 | -97.5014
4.2832 | -97.5655
4.2869 | -97.6297
4.2907 | -97.6772
4.2944 | -97.7109
4.2982 | -97.7398
4.3019 | -97.7583
4.3057 | -97.7793
4.3094 | -97.8188
4.3132 | -97.8675
4.3169 | -97.911
4.3207 | -97.9615
4.3244 | -98.0165
4.3282 | -98.062
4.3319 | -98.1126
4.3357 | -98.1538
4.3394 | -98.1718
4.3432 | -98.1764
4.3469 | -98.1629
4.3507 | -98.1195
4.3544 | -98.046
4.3582 | -97.952
4.3619 | -97.8503
4.3657 | -97.7497
4.3694 | -97.6651
4.3732 | -97.5992
4.3769 | -97.5497
4.3807 | -97.5293
4.3844 | -97.5356
4.3882 | -97.5546
4.3919 | -97.5874
4.3957 | -97.6256
4.3994 | -97.6547
4.4032 | -97.6674
4.4069 | -97.6655
4.4107 | -97.6497
4.4144 | -97.6191
4.4182 | -97.5903
4.4219 | -97.5731
4.4257 | -97.5759
4.4294 | -97.6059
4.4332 | -97.6533
4.4369 | -97.7166
4.4407 | -97.785
4.4444 | -97.8533
4.4482 | -97.9233
4.4519 | -97.9977
4.4557 | -98.0779
4.4594 | -98.1595
4.4632 | -98.2385
4.4669 | -98.3113
4.4707 | -98.3578
4.4744 | -98.3683
4.4782 | -98.3398
4.482 | -98.267
4.4857 | -98.1499
4.4895 | -97.9993
4.4932 | -97.8425
4.497 | -97.6909
4.5007 | -97.5473
4.5045 | -97.4288
4.5082 | -97.3432
4.512 | -97.2941
4.5157 | -97.2768
4.5195 | -97.2889
4.5232 | -97.3269
4.527 | -97.3913
4.5307 | -97.4787
4.5345 | -97.5849
4.5382 | -97.7049
4.542 | -97.8303
4.5457 | -97.9468
4.5495 | -98.044
4.5532 | -98.1132
4.557 | -98.1567
4.5607 | -98.1788
4.5645 | -98.1781
4.5682 | -98.1713
4.572 | -98.17
4.5757 | -98.1716
4.5795 | -98.1783
4.5832 | -98.194
4.587 | -98.2114
4.5907 | -98.2193
4.5945 | -98.2179
4.5982 | -98.2
4.602 | -98.172
4.6057 | -98.1417
4.6095 | -98.1187
4.6132 | -98.1076
4.617 | -98.1101
4.6207 | -98.126
4.6245 | -98.1554
4.6282 | -98.19
4.632 | -98.2254
4.6357 | -98.2582
4.6395 | -98.2867
4.6432 | -98.315
4.647 | -98.3278
4.6507 | -98.3185
4.6545 | -98.2814
4.6582 | -98.2304
4.662 | -98.1748
4.6657 | -98.1179
4.6695 | -98.0753
4.6732 | -98.0554
4.677 | -98.0561
4.6807 | -98.0835
4.6845 | -98.1368
4.6882 | -98.2101
4.692 | -98.2964
4.6957 | -98.3938
4.6995 | -98.4882
4.7032 | -98.5725
4.707 | -98.6452
4.7107 | -98.6981
4.7145 | -98.7312
4.7182 | -98.7627
4.722 | -98.7895
4.7257 | -98.8171
4.7295 | -98.8494
4.7332 | -98.8877
4.737 | -98.9322
4.7407 | -98.9637
4.7445 | -98.9742
4.7482 | -98.9644
4.752 | -98.9307
4.7557 | -98.8712
4.7595 | -98.7933
4.7632 | -98.7105
4.767 | -98.6224
4.7707 | -98.5426
4.7745 | -98.4779
4.7782 | -98.4236
4.782 | -98.3834
4.7857 | -98.3516
4.7895 | -98.3238
4.7932 | -98.3047
4.797 | -98.2978
4.8007 | -98.3012
4.8045 | -98.3181
4.8083 | -98.3468
4.812 | -98.3967
4.8158 | -98.4599
4.8195 | -98.5316
4.8233 | -98.609
4.827 | -98.6884
4.8308 | -98.7623
4.8345 | -98.8296
4.8383 | -98.8967
4.842 | -98.9489
4.8458 | -98.9794
4.8495 | -98.9886
4.8533 | -98.9773
4.857 | -98.9488
4.8608 | -98.9078
4.8645 | -98.859
4.8683 | -98.8367
4.872 | -98.815
4.8758 | -98.8021
4.8795 | -98.8078
4.8833 | -98.8338
4.887 | -98.8682
4.8908 | -98.9067
4.8945 | -98.9389
4.8983 | -98.961
4.902 | -98.9725
4.9058 | -98.9701
4.9095 | -98.9585
4.9133 | -98.9382
4.917 | -98.9206
4.9208 | -98.9058
4.9245 | -98.9014
4.9283 | -98.9086
4.932 | -98.9206
4.9358 | -98.9361
4.9395 | -98.9505
4.9433 | -98.9649
4.947 | -98.9709
4.9508 | -98.9771
4.9545 | -98.9848
4.9583 | -98.9804
4.962 | -98.9669
4.9658 | -98.9557
4.9695 | -98.9489
4.9733 | -98.9469
4.977 | -98.9475
4.9808 | -98.9489
4.9845 | -98.9458
4.9883 | -98.9405
4.992 | -98.9337
4.9958 | -98.9277
4.9995 | -98.9232
5.0033 | -98.9173
5.007 | -98.9145
5.0108 | -98.9232
5.0145 | -98.9435
5.0183 | -98.9736
5.022 | -99.0068
5.0258 | -99.0478
5.0295 | -99.0954
5.0333 | -99.1491
5.037 | -99.214
5.0408 | -99.2834
5.0445 | -99.3426
5.0483 | -99.3932
5.052 | -99.4438
5.0558 | -99.4806
5.0595 | -99.4971
5.0633 | -99.4981
5.067 | -99.4738
5.0708 | -99.4296
5.0745 | -99.3627
5.0783 | -99.2777
5.082 | -99.1855
5.0858 | -99.094
5.0895 | -99.0136
5.0933 | -98.95
5.097 | -98.9093
5.1008 | -98.8897
5.1045 | -98.889
5.1083 | -98.9017
5.112 | -98.9273
5.1158 | -98.9604
5.1196 | -99.002
5.1233 | -99.0571
5.1271 | -99.1179
5.1308 | -99.1786
5.1346 | -99.2492
5.1383 | -99.3221
5.1421 | -99.3907
5.1458 | -99.4481
5.1496 | -99.4876
5.1533 | -99.5123
5.1571 | -99.5259
5.1608 | -99.5198
5.1646 | -99.4924
5.1683 | -99.4421
5.1721 | -99.3714
5.1758 | -99.2989
5.1796 | -99.2354
5.1833 | -99.1876
5.1871 | -99.163
5.1908 | -99.1613
5.1946 | -99.1825
5.1983 | -99.2203
5.2021 | -99.2627
5.2058 | -99.3026
5.2096 | -99.3309
5.2133 | -99.3448
5.2171 | -99.3503
5.2208 | -99.349
5.2246 | -99.3498
5.2283 | -99.3501
5.2321 | -99.3482
5.2358 | -99.3445
5.2396 | -99.3446
5.2433 | -99.3451
5.2471 | -99.35
5.2508 | -99.353
5.2546 | -99.3478
5.2583 | -99.3471
5.2621 | -99.3468
5.2658 | -99.3431
5.2696 | -99.3388
5.2733 | -99.3288
5.2771 | -99.3081
5.2808 | -99.282
5.2846 | -99.2491
5.2883 | -99.2126
5.2921 | -99.1669
5.2958 | -99.1095
5.2996 | -99.0469
5.3033 | -98.9829
5.3071 | -98.9283
5.3108 | -98.8854
5.3146 | -98.8596
5.3183 | -98.8542
5.3221 | -98.878
5.3258 | -98.9242
5.3296 | -98.9916
5.3333 | -99.0676
5.3371 | -99.1414
5.3408 | -99.2166
5.3446 | -99.2951
5.3483 | -99.3722
5.3521 | -99.4474
5.3558 | -99.5163
5.3596 | -99.5763
5.3633 | -99.6234
5.3671 | -99.649
5.3708 | -99.6529
5.3746 | -99.64
5.3783 | -99.6087
5.3821 | -99.5668
5.3858 | -99.5204
5.3896 | -99.4855
5.3933 | -99.4704
5.3971 | -99.4808
5.4008 | -99.5161
5.4046 | -99.5784
5.4083 | -99.6781
5.4121 | -99.7536
5.4158 | -99.8346
5.4196 | -99.906
5.4233 | -99.9629
5.4271 | -100.0038
5.4308 | -100.0283
5.4346 | -100.0407
5.4383 | -100.0461
5.4421 | -100.0411
5.4459 | -100.0331
5.4496 | -100.0246
5.4534 | -100.0232
5.4571 | -100.0328
5.4609 | -100.0465
5.4646 | -100.0684
5.4684 | -100.0969
5.4721 | -100.1238
5.4759 | -100.1471
5.4796 | -100.1648
5.4834 | -100.1653
5.4871 | -100.1495
5.4909 | -100.1213
5.4946 | -100.0774
5.4984 | -100.0208
5.5021 | -99.9572
5.5059 | -99.8913
5.5096 | -99.8232
5.5134 | -99.7554
5.5171 | -99.6954
5.5209 | -99.644
5.5246 | -99.6024
5.5284 | -99.5742
5.5321 | -99.5634
5.5359 | -99.5634
5.5396 | -99.5684
5.5434 | -99.5773
5.5471 | -99.5944
5.5509 | -99.6166
5.5546 | -99.6405
5.5584 | -99.6662
5.5621 | -99.6948
5.5659 | -99.7259
5.5696 | -99.7633
5.5734 | -99.8024
5.5771 | -99.8436
5.5809 | -99.8904
5.5846 | -99.9398
5.5884 | -99.9914
5.5921 | -100.0485
5.5959 | -100.0969
5.5996 | -100.1326
5.6034 | -100.1587
5.6071 | -100.1748
5.6109 | -100.1792
5.6146 | -100.1725
5.6184 | -100.1534
5.6221 | -100.1314
5.6259 | -100.1015
5.6296 | -100.0658
5.6334 | -100.038
5.6371 | -100.0175
5.6409 | -100.008
5.6446 | -100.0088
5.6484 | -100.0203
5.6521 | -100.0517
5.6559 | -100.1
5.6596 | -100.16
5.6634 | -100.2299
5.6671 | -100.3019
5.6709 | -100.3717
5.6746 | -100.439
5.6784 | -100.4948
5.6821 | -100.548
5.6859 | -100.5853
5.6896 | -100.6018
5.6934 | -100.5986
5.6971 | -100.5753
5.7009 | -100.5408
5.7046 | -100.4941
5.7084 | -100.4368
5.7121 | -100.3665
5.7159 | -100.2865
5.7196 | -100.2011
5.7234 | -100.1198
5.7271 | -100.0511
5.7309 | -99.9956
5.7346 | -99.9616
5.7384 | -99.9522
5.7421 | -99.9659
5.7459 | -100.0065
5.7496 | -100.0681
5.7534 | -100.1486
5.7572 | -100.2408
5.7609 | -100.335
5.7647 | -100.4283
5.7684 | -100.5137
5.7722 | -100.5975
5.7759 | -100.6608
5.7797 | -100.7018
5.7834 | -100.7219
5.7872 | -100.7105
5.7909 | -100.6746
5.7947 | -100.6231
5.7984 | -100.5623
5.8022 | -100.4933
5.8059 | -100.4195
5.8097 | -100.3511
5.8134 | -100.2987
5.8172 | -100.2619
5.8209 | -100.2319
5.8247 | -100.2069
5.8284 | -100.1942
5.8322 | -100.1864
5.8359 | -100.1778
5.8397 | -100.1686
5.8434 | -100.1582
5.8472 | -100.1472
5.8509 | -100.1386
5.8547 | -100.1419
5.8584 | -100.1528
5.8622 | -100.18
5.8659 | -100.2241
5.8697 | -100.2812
5.8734 | -100.3546
5.8772 | -100.4403
5.8809 | -100.5285
5.8847 | -100.6177
5.8884 | -100.7072
5.8922 | -100.7875
5.8959 | -100.8484
5.8997 | -100.8909
5.9034 | -100.9273
5.9072 | -100.9517
5.9109 | -100.9547
5.9147 | -100.9413
5.9184 | -100.9185
5.9222 | -100.8922
5.9259 | -100.8584
5.9297 | -100.8244
5.9334 | -100.7843
5.9372 | -100.7354
5.9409 | -100.6815
5.9447 | -100.6312
5.9484 | -100.5756
5.9522 | -100.5186
5.9559 | -100.4686
5.9597 | -100.4284
5.9634 | -100.398
5.9672 | -100.3765
5.9709 | -100.3698
5.9747 | -100.3807
5.9784 | -100.4052
5.9822 | -100.4468
5.9859 | -100.5068
5.9897 | -100.5749
5.9934 | -100.6481
5.9972 | -100.728
6.0009 | -100.8136
6.0047 | -100.8952
6.0084 | -100.9705
6.0122 | -101.0391
6.0159 | -101.104
6.0197 | -101.1521
6.0234 | -101.1843
6.0272 | -101.1982
6.0309 | -101.1856
6.0347 | -101.1489
6.0384 | -101.1006
6.0422 | -101.052
6.0459 | -101.0035
6.0497 | -100.9486
6.0534 | -100.8954
6.0572 | -100.8508
6.0609 | -100.8202
6.0647 | -100.7972
6.0684 | -100.779
6.0722 | -100.7685
6.0759 | -100.7676
6.0797 | -100.7704
6.0835 | -100.7767
6.0872 | -100.7895
6.091 | -100.8072
6.0947 | -100.8273
6.0985 | -100.8385
6.1022 | -100.8433
6.106 | -100.8386
6.1097 | -100.8232
6.1135 | -100.8025
6.1172 | -100.7788
6.121 | -100.7534
6.1247 | -100.7312
6.1285 | -100.7147
6.1322 | -100.7074
6.136 | -100.7046
6.1397 | -100.7091
6.1435 | -100.7233
6.1472 | -100.7477
6.151 | -100.7755
6.1547 | -100.8008
6.1585 | -100.8229
6.1622 | -100.8345
6.166 | -100.8385
6.1697 | -100.8352
6.1735 | -100.8297
6.1772 | -100.8245
6.181 | -100.8129
6.1847 | -100.8039
6.1885 | -100.7987
6.1922 | -100.8011
6.196 | -100.7938
6.1997 | -100.7803
6.2035 | -100.7615
6.2072 | -100.7395
6.211 | -100.7131
6.2147 | -100.6837
6.2185 | -100.653
6.2222 | -100.6256
6.226 | -100.5923
6.2297 | -100.5603
6.2335 | -100.54
6.2372 | -100.5244
6.241 | -100.511
6.2447 | -100.5045
6.2485 | -100.507
6.2522 | -100.5192
6.256 | -100.5426
6.2597 | -100.5733
6.2635 | -100.6128
6.2672 | -100.6634
6.271 | -100.725
6.2747 | -100.8006
6.2785 | -100.8838
6.2822 | -100.9744
6.286 | -101.0699
6.2897 | -101.1641
6.2935 | -101.245
6.2972 | -101.3102
6.301 | -101.356
6.3047 | -101.3776
6.3085 | -101.3751
6.3122 | -101.3494
6.316 | -101.3023
6.3197 | -101.249
6.3235 | -101.1923
6.3272 | -101.1286
6.331 | -101.0754
6.3347 | -101.0369
6.3385 | -101.0122
6.3422 | -101.0024
6.346 | -101.0048
6.3497 | -101.0193
6.3535 | -101.0448
6.3572 | -101.0748
6.361 | -101.1031
6.3647 | -101.1288
6.3685 | -101.1476
6.3722 | -101.1585
6.376 | -101.1578
6.3797 | -101.146
6.3835 | -101.1325
6.3872 | -101.1111
6.391 | -101.0806
6.3947 | -101.042
6.3985 | -101.0036
6.4023 | -100.963
6.406 | -100.9234
6.4098 | -100.8863
6.4135 | -100.8467
6.4173 | -100.8056
6.421 | -100.7722
6.4248 | -100.7402
6.4285 | -100.7129
6.4323 | -100.6949
6.436 | -100.6842
6.4398 | -100.6775
6.4435 | -100.6743
6.4473 | -100.676
6.451 | -100.688
6.4548 | -100.7066
6.4585 | -100.7254
6.4623 | -100.7496
6.466 | -100.779
6.4698 | -100.8126
6.4735 | -100.8487
6.4773 | -100.8847
6.481 | -100.9292
6.4848 | -100.9822
6.4885 | -101.0436
6.4923 | -101.1078
6.496 | -101.1684
6.4998 | -101.2297
6.5035 | -101.2874
6.5073 | -101.3384
6.511 | -101.3853
6.5148 | -101.4241
6.5185 | -101.4442
6.5223 | -101.4556
6.526 | -101.4564
6.5298 | -101.4424
6.5335 | -101.4198
6.5373 | -101.3939
6.541 | -101.3682
6.5448 | -101.3377
6.5485 | -101.2996
6.5523 | -101.2579
6.556 | -101.2193
6.5598 | -101.18
6.5635 | -101.1365
6.5673 | -101.091
6.571 | -101.0498
6.5748 | -101.0169
6.5785 | -100.9853
6.5823 | -100.959
6.586 | -100.942
6.5898 | -100.9368
6.5935 | -100.9383
6.5973 | -100.944
6.601 | -100.9564
6.6048 | -100.9758
6.6085 | -101.0041
6.6123 | -101.0357
6.616 | -101.0729
6.6198 | -101.1214
6.6235 | -101.175
6.6273 | -101.2305
6.631 | -101.2898
6.6348 | -101.3505
6.6385 | -101.4069
6.6423 | -101.4497
6.646 | -101.4805
6.6498 | -101.5045
6.6535 | -101.5216
6.6573 | -101.5232
6.661 | -101.5158
6.6648 | -101.5036
6.6685 | -101.4821
6.6723 | -101.4533
6.676 | -101.418
6.6798 | -101.3788
6.6835 | -101.338
6.6873 | -101.3021
6.691 | -101.2685
6.6948 | -101.2394
6.6985 | -101.2148
6.7023 | -101.1967
6.706 | -101.1792
6.7098 | -101.1682
6.7135 | -101.1669
6.7173 | -101.1661
6.721 | -101.1634
6.7248 | -101.161
6.7286 | -101.1658
6.7323 | -101.1708
6.7361 | -101.1725
6.7398 | -101.1769
6.7436 | -101.1867
6.7473 | -101.1997
6.7511 | -101.2115
6.7548 | -101.2231
6.7586 | -101.2381
6.7623 | -101.257
6.7661 | -101.2799
6.7698 | -101.3117
6.7736 | -101.3432
6.7773 | -101.381
6.7811 | -101.4291
6.7848 | -101.4752
6.7886 | -101.5235
6.7923 | -101.5841
6.7961 | -101.6485
6.7998 | -101.7079
6.8036 | -101.7636
6.8073 | -101.8189
6.8111 | -101.8583
6.8148 | -101.8891
6.8186 | -101.9206
6.8223 | -101.9491
6.8261 | -101.9736
6.8298 | -101.9949
6.8336 | -102.0179
6.8373 | -102.0342
6.8411 | -102.0458
6.8448 | -102.0556
6.8486 | -102.0589
6.8523 | -102.0555
6.8561 | -102.0452
6.8598 | -102.0226
6.8636 | -101.9837
6.8673 | -101.9411
6.8711 | -101.8947
6.8748 | -101.8523
6.8786 | -101.8078
6.8823 | -101.7563
6.8861 | -101.7106
6.8898 | -101.6646
6.8936 | -101.6145
6.8973 | -101.5722
6.9011 | -101.5409
6.9048 | -101.5143
6.9086 | -101.4849
6.9123 | -101.4621
6.9161 | -101.4444
6.9198 | -101.4331
6.9236 | -101.4263
6.9273 | -101.4233
6.9311 | -101.4247
6.9348 | -101.4325
6.9386 | -101.4458
6.9423 | -101.4648
6.9461 | -101.4882
6.9498 | -101.5162
6.9536 | -101.546
6.9573 | -101.5786
6.9611 | -101.6211
6.9648 | -101.6691
6.9686 | -101.712
6.9723 | -101.75
6.9761 | -101.7874
6.9798 | -101.8217
6.9836 | -101.8527
6.9873 | -101.8759
6.9911 | -101.8923
6.9948 | -101.9044
6.9986 | -101.9123
7.0023 | -101.9213
7.0061 | -101.9315
7.0098 | -101.9439
7.0136 | -101.9573
7.0173 | -101.9702
7.0211 | -101.9806
7.0248 | -101.9964
7.0286 | -102.0121
7.0323 | -102.0224
7.0361 | -102.0275
7.0399 | -102.0332
7.0436 | -102.0433
7.0474 | -102.045
7.0511 | -102.0414
7.0549 | -102.0352
7.0586 | -102.0233
7.0624 | -102.0084
7.0661 | -101.9867
7.0699 | -101.9625
7.0736 | -101.9358
7.0774 | -101.9078
7.0811 | -101.8799
7.0849 | -101.8538
7.0886 | -101.8343
7.0924 | -101.8182
7.0961 | -101.805
7.0999 | -101.8021
7.1036 | -101.8063
7.1074 | -101.8185
7.1111 | -101.8413
7.1149 | -101.8788
7.1186 | -101.9237
7.1224 | -101.9705
7.1261 | -102.0214
7.1299 | -102.0747
7.1336 | -102.1333
7.1374 | -102.1919
7.1411 | -102.2471
7.1449 | -102.2982
7.1486 | -102.3371
7.1524 | -102.3611
7.1561 | -102.3749
7.1599 | -102.3782
7.1636 | -102.3741
7.1674 | -102.3614
7.1711 | -102.3371
7.1749 | -102.3059
7.1786 | -102.2748
7.1824 | -102.2413
7.1861 | -102.204
7.1899 | -102.1635
7.1936 | -102.1288
7.1974 | -102.0963
7.2011 | -102.0645
7.2049 | -102.0334
7.2086 | -102.0019
7.2124 | -101.9753
7.2161 | -101.9498
7.2199 | -101.9283
7.2236 | -101.9071
7.2274 | -101.8865
7.2311 | -101.8703
7.2349 | -101.8565
7.2386 | -101.8363
7.2424 | -101.8137
7.2461 | -101.7919
7.2499 | -101.7682
7.2536 | -101.7402
7.2574 | -101.7155
7.2611 | -101.6927
7.2649 | -101.6652
7.2686 | -101.6359
7.2724 | -101.608
7.2761 | -101.5869
7.2799 | -101.5707
7.2836 | -101.5606
7.2874 | -101.5552
7.2911 | -101.5635
7.2949 | -101.586
7.2986 | -101.6211
7.3024 | -101.6642
7.3061 | -101.7156
7.3099 | -101.772
7.3136 | -101.8372
7.3174 | -101.9056
7.3211 | -101.9742
7.3249 | -102.0397
7.3286 | -102.0985
7.3324 | -102.1564
7.3361 | -102.2102
7.3399 | -102.2563
7.3436 | -102.2889
7.3474 | -102.3139
7.3511 | -102.3263
7.3549 | -102.324
7.3586 | -102.3184
7.3624 | -102.3103
7.3662 | -102.2919
7.3699 | -102.2697
7.3737 | -102.2509
7.3774 | -102.2286
7.3812 | -102.2042
7.3849 | -102.1782
7.3887 | -102.1561
7.3924 | -102.1372
7.3962 | -102.1238
7.3999 | -102.1094
7.4037 | -102.0953
7.4074 | -102.0831
7.4112 | -102.0802
7.4149 | -102.0819
7.4187 | -102.0876
7.4224 | -102.0942
7.4262 | -102.1055
7.4299 | -102.1175
7.4337 | -102.1355
7.4374 | -102.1596
7.4412 | -102.1834
7.4449 | -102.2078
7.4487 | -102.2306
7.4524 | -102.2551
7.4562 | -102.2772
7.4599 | -102.2999
7.4637 | -102.3233
7.4674 | -102.3412
7.4712 | -102.3583
7.4749 | -102.3801
7.4787 | -102.4039
7.4824 | -102.4267
7.4862 | -102.4503
7.4899 | -102.4683
7.4937 | -102.4879
7.4974 | -102.505
7.5012 | -102.5208
7.5049 | -102.5321
7.5087 | -102.5447
7.5124 | -102.5578
7.5162 | -102.5643
7.5199 | -102.5597
7.5237 | -102.5536
7.5274 | -102.5461
7.5312 | -102.5323
7.5349 | -102.5135
7.5387 | -102.4948
7.5424 | -102.4734
7.5462 | -102.4532
7.5499 | -102.4426
7.5537 | -102.4412
7.5574 | -102.4463
7.5612 | -102.4584
7.5649 | -102.4785
7.5687 | -102.4949
7.5724 | -102.5014
7.5762 | -102.5116
7.5799 | -102.5217
7.5837 | -102.5255
7.5874 | -102.5201
7.5912 | -102.5129
7.5949 | -102.4997
7.5987 | -102.4848
7.6024 | -102.4727
7.6062 | -102.4549
7.6099 | -102.442
7.6137 | -102.4376
7.6174 | -102.4427
7.6212 | -102.4518
7.6249 | -102.4659
7.6287 | -102.4846
7.6324 | -102.5081
7.6362 | -102.5358
7.6399 | -102.5657
7.6437 | -102.5907
7.6474 | -102.6097
7.6512 | -102.6236
7.6549 | -102.6307
7.6587 | -102.6244
7.6624 | -102.6046
7.6662 | -102.5797
7.6699 | -102.5477
7.6737 | -102.5046
7.6775 | -102.4534
7.6812 | -102.4025
7.685 | -102.3458
7.6887 | -102.2834
7.6925 | -102.2233
7.6962 | -102.1668
7.7 | -102.1133
7.7037 | -102.0672
7.7075 | -102.0261
7.7112 | -101.9899
7.715 | -101.959
7.7187 | -101.9341
7.7225 | -101.9162
7.7262 | -101.9003
7.73 | -101.8916
7.7337 | -101.8921
7.7375 | -101.9019
7.7412 | -101.9211
7.745 | -101.9507
7.7487 | -101.9838
7.7525 | -102.0198
7.7562 | -102.0613
7.76 | -102.1139
7.7637 | -102.1689
7.7675 | -102.2228
7.7712 | -102.2758
7.775 | -102.3319
7.7787 | -102.3911
7.7825 | -102.447
7.7862 | -102.5
7.79 | -102.5496
7.7937 | -102.5958
7.7975 | -102.6402
7.8012 | -102.6873
7.805 | -102.728
7.8087 | -102.76
7.8125 | -102.788
7.8162 | -102.816
7.82 | -102.8379
7.8237 | -102.8558
7.8275 | -102.8718
7.8312 | -102.8862
7.835 | -102.9023
7.8387 | -102.9205
7.8425 | -102.9401
7.8462 | -102.949
7.85 | -102.9474
7.8537 | -102.9452
7.8575 | -102.9382
7.8612 | -102.9219
7.865 | -102.8976
7.8687 | -102.8594
7.8725 | -102.8116
7.8762 | -102.7554
7.88 | -102.6949
7.8837 | -102.6343
7.8875 | -102.5755
7.8912 | -102.521
7.895 | -102.4737
7.8987 | -102.4375
7.9025 | -102.4169
7.9062 | -102.4109
7.91 | -102.4186
7.9137 | -102.4347
7.9175 | -102.4596
7.9212 | -102.4959
7.925 | -102.5373
7.9287 | -102.5797
7.9325 | -102.6235
7.9362 | -102.6649
7.94 | -102.6986
7.9437 | -102.7265
7.9475 | -102.7416
7.9512 | -102.742
7.955 | -102.7346
7.9587 | -102.7184
7.9625 | -102.7004
7.9662 | -102.6819
7.97 | -102.6551
7.9737 | -102.6199
7.9775 | -102.5863
7.9812 | -102.556
7.985 | -102.528
7.9887 | -102.4991
7.9925 | -102.471
7.9962 | -102.4422
8 | -102.4165
8.0038 | -102.3936
8.0075 | -102.3747
8.0113 | -102.3619
8.015 | -102.3573
8.0188 | -102.3626
8.0225 | -102.378
8.0263 | -102.3987
8.03 | -102.4297
8.0338 | -102.4681
8.0375 | -102.5122
8.0413 | -102.5602
8.045 | -102.6096
8.0488 | -102.6573
8.0525 | -102.7063
8.0563 | -102.7515
8.06 | -102.7964
8.0638 | -102.8444
8.0675 | -102.8957
8.0713 | -102.9436
8.075 | -102.9955
8.0788 | -103.0467
8.0825 | -103.0959
8.0863 | -103.1477
8.09 | -103.2032
8.0938 | -103.252
8.0975 | -103.2911
8.1013 | -103.3207
8.105 | -103.3346
8.1088 | -103.337
8.1125 | -103.3322
8.1163 | -103.2867
8.12 | -103.2649
8.1238 | -103.2348
8.1275 | -103.1997
8.1313 | -103.1698
8.135 | -103.1409
8.1388 | -103.112
8.1425 | -103.083
8.1463 | -103.064
8.15 | -103.0535
8.1538 | -103.0479
8.1575 | -103.0424
8.1613 | -103.0401
8.165 | -103.0427
8.1688 | -103.0486
8.1725 | -103.0525
8.1763 | -103.0541
8.18 | -103.0491
8.1838 | -103.0414
8.1875 | -103.0319
8.1913 | -103.0253
8.195 | -103.0171
8.1988 | -103.0032
8.2025 | -102.987
8.2063 | -102.9658
8.21 | -102.9443
8.2138 | -102.9279
8.2175 | -102.9177
8.2213 | -102.9127
8.225 | -102.9125
8.2288 | -102.9135
8.2325 | -102.9131
8.2363 | -102.9157
8.24 | -102.9235
8.2438 | -102.9352
8.2475 | -102.9503
8.2513 | -102.9657
8.255 | -102.9766
8.2588 | -102.9921
8.2625 | -103.0135
8.2663 | -103.0406
8.27 | -103.0698
8.2738 | -103.1005
8.2775 | -103.1298
8.2813 | -103.1647
8.285 | -103.203
8.2888 | -103.2388
8.2925 | -103.2751
8.2963 | -103.3137
8.3 | -103.352
8.3038 | -103.3869
8.3075 | -103.414
8.3113 | -103.4401
8.315 | -103.4582
8.3188 | -103.4718
8.3225 | -103.4858
8.3263 | -103.4943
8.3301 | -103.5025
8.3338 | -103.5057
8.3376 | -103.4989
8.3413 | -103.4803
8.3451 | -103.4587
8.3488 | -103.4367
8.3526 | -103.4123
8.3563 | -103.3901
8.3601 | -103.3645
8.3638 | -103.3343
8.3676 | -103.3022
8.3713 | -103.2777
8.3751 | -103.2592
8.3788 | -103.243
8.3826 | -103.2299
8.3863 | -103.2175
8.3901 | -103.2061
8.3938 | -103.1977
8.3976 | -103.1905
8.4013 | -103.1838
8.4051 | -103.179
8.4088 | -103.1706
8.4126 | -103.1586
8.4163 | -103.1431
8.4201 | -103.1244
8.4238 | -103.1047
8.4276 | -103.0877
8.4313 | -103.0703
8.4351 | -103.0453
8.4388 | -103.0186
8.4426 | -102.9946
8.4463 | -102.9709
8.4501 | -102.9446
8.4538 | -102.9163
8.4576 | -102.8873
8.4613 | -102.8631
8.4651 | -102.8469
8.4688 | -102.8271
8.4726 | -102.8095
8.4763 | -102.7947
8.4801 | -102.7861
8.4838 | -102.7832
8.4876 | -102.7879
8.4913 | -102.7973
8.4951 | -102.8066
8.4988 | -102.8215
8.5026 | -102.8443
8.5063 | -102.8747
8.5101 | -102.9089
8.5138 | -102.9496
8.5176 | -102.9962
8.5213 | -103.0477
8.5251 | -103.1035
8.5288 | -103.1575
8.5326 | -103.216
8.5363 | -103.2819
8.5401 | -103.3479
8.5438 | -103.4101
8.5476 | -103.4698
8.5513 | -103.5211
8.5551 | -103.5687
8.5588 | -103.6103
8.5626 | -103.6441
8.5663 | -103.6641
8.5701 | -103.6741
8.5738 | -103.6769
8.5776 | -103.668
8.5813 | -103.6496
8.5851 | -103.6252
8.5888 | -103.5974
8.5926 | -103.5701
8.5963 | -103.5408
8.6001 | -103.5105
8.6038 | -103.4808
8.6076 | -103.4563
8.6113 | -103.436
8.6151 | -103.4166
8.6188 | -103.3963
8.6226 | -103.3685
8.6263 | -103.3382
8.6301 | -103.3062
8.6338 | -103.2756
8.6376 | -103.2438
8.6413 | -103.2118
8.6451 | -103.1805
8.6489 | -103.1527
8.6526 | -103.1208
8.6564 | -103.0978
8.6601 | -103.0738
8.6639 | -103.0511
8.6676 | -103.0352
8.6714 | -103.0252
8.6751 | -103.0178
8.6789 | -103.0173
8.6826 | -103.0225
8.6864 | -103.0321
8.6901 | -103.0462
8.6939 | -103.0617
8.6976 | -103.078
8.7014 | -103.0943
8.7051 | -103.1115
8.7089 | -103.13
8.7126 | -103.143
8.7164 | -103.1469
8.7201 | -103.1489
8.7239 | -103.1502
8.7276 | -103.1539
8.7314 | -103.1559
8.7351 | -103.16
8.7389 | -103.167
8.7426 | -103.1786
8.7464 | -103.1967
8.7501 | -103.2231
8.7539 | -103.2552
8.7576 | -103.2877
8.7614 | -103.3241
8.7651 | -103.3625
8.7689 | -103.404
8.7726 | -103.4478
8.7764 | -103.4939
8.7801 | -103.5422
8.7839 | -103.5865
8.7876 | -103.6328
8.7914 | -103.676
8.7951 | -103.7137
8.7989 | -103.7479
8.8026 | -103.7769
8.8064 | -103.7999
8.8101 | -103.8159
8.8139 | -103.8257
8.8176 | -103.8403
8.8214 | -103.8491
8.8251 | -103.8556
8.8289 | -103.8566
8.8326 | -103.8508
8.8364 | -103.8394
8.8401 | -103.8245
8.8439 | -103.8078
8.8476 | -103.7928
8.8514 | -103.7727
8.8551 | -103.7477
8.8589 | -103.7201
8.8626 | -103.6926
8.8664 | -103.6598
8.8701 | -103.6304
8.8739 | -103.6046
8.8776 | -103.5818
8.8814 | -103.5601
8.8851 | -103.5362
8.8889 | -103.5165
8.8926 | -103.4995
8.8964 | -103.4875
8.9001 | -103.471
8.9039 | -103.4577
8.9076 | -103.4465
8.9114 | -103.4358
8.9151 | -103.4237
8.9189 | -103.4089
8.9226 | -103.3953
8.9264 | -103.3822
8.9301 | -103.3693
8.9339 | -103.3625
8.9376 | -103.3559
8.9414 | -103.3456
8.9451 | -103.3359
8.9489 | -103.3219
8.9526 | -103.3115
8.9564 | -103.3031
8.9602 | -103.2959
8.9639 | -103.2908
8.9677 | -103.2826
8.9714 | -103.2718
8.9752 | -103.2621
8.9789 | -103.2492
8.9827 | -103.2379
8.9864 | -103.2274
8.9902 | -103.2148
8.9939 | -103.2002
8.9977 | -103.1897
9.0014 | -103.1795
9.0052 | -103.1683
9.0089 | -103.1547
9.0127 | -103.1435
9.0164 | -103.1371
9.0202 | -103.1294
9.0239 | -103.1228
9.0277 | -103.1177
9.0314 | -103.1202
9.0352 | -103.1261
9.0389 | -103.1353
9.0427 | -103.1473
9.0464 | -103.1635
9.0502 | -103.1835
9.0539 | -103.2061
9.0577 | -103.237
9.0614 | -103.2692
9.0652 | -103.2952
9.0689 | -103.329
9.0727 | -103.3692
9.0764 | -103.4089
9.0802 | -103.4452
9.0839 | -103.4778
9.0877 | -103.502
9.0914 | -103.5208
9.0952 | -103.5341
9.0989 | -103.5448
9.1027 | -103.5466
9.1064 | -103.5349
9.1102 | -103.5155
9.1139 | -103.4906
9.1177 | -103.4607
9.1214 | -103.427
9.1252 | -103.3914
9.1289 | -103.3538
9.1327 | -103.3148
9.1364 | -103.2796
9.1402 | -103.2538
9.1439 | -103.2304
9.1477 | -103.2104
9.1514 | -103.2009
9.1552 | -103.2004
9.1589 | -103.2078
9.1627 | -103.2203
9.1664 | -103.2382
9.1702 | -103.2602
9.1739 | -103.2873
9.1777 | -103.3191
9.1814 | -103.3523
9.1852 | -103.3867
9.1889 | -103.4189
9.1927 | -103.4521
9.1964 | -103.4693
9.2002 | -103.4945
9.2039 | -103.5096
9.2077 | -103.5196
9.2114 | -103.5301
9.2152 | -103.5407
9.2189 | -103.5477
9.2227 | -103.5531
9.2264 | -103.5563
9.2302 | -103.5561
9.2339 | -103.554
9.2377 | -103.5517
9.2414 | -103.5543
9.2452 | -103.5564
9.2489 | -103.551
9.2527 | -103.5484
9.2564 | -103.5534
9.2602 | -103.559
9.2639 | -103.5617
9.2677 | -103.5633
9.2714 | -103.569
9.2752 | -103.5733
9.279 | -103.578
9.2827 | -103.5814
9.2865 | -103.5849
9.2902 | -103.5832
9.294 | -103.5802
9.2977 | -103.581
9.3015 | -103.5801
9.3052 | -103.5751
9.309 | -103.575
9.3127 | -103.578
9.3165 | -103.5822
9.3202 | -103.5856
9.324 | -103.5932
9.3277 | -103.6023
9.3315 | -103.6142
9.3352 | -103.6283
9.339 | -103.6413
9.3427 | -103.654
9.3465 | -103.6683
9.3502 | -103.6763
9.354 | -103.685
9.3577 | -103.6912
9.3615 | -103.6997
9.3652 | -103.7091
9.369 | -103.718
9.3727 | -103.7278
9.3765 | -103.7412
9.3802 | -103.7526
9.384 | -103.7691
9.3877 | -103.7903
9.3915 | -103.8084
9.3952 | -103.8244
9.399 | -103.8439
9.4027 | -103.8677
9.4065 | -103.8926
9.4102 | -103.9122
9.414 | -103.9331
9.4177 | -103.9577
9.4215 | -103.9884
9.4252 | -104.0163
9.429 | -104.0463
9.4327 | -104.0775
9.4365 | -104.1028
9.4402 | -104.1292
9.444 | -104.1606
9.4477 | -104.189
9.4515 | -104.2089
9.4552 | -104.2222
9.459 | -104.2308
9.4627 | -104.2342
9.4665 | -104.2359
9.4702 | -104.2351
9.474 | -104.2335
9.4777 | -104.2359
9.4815 | -104.2348
9.4852 | -104.2266
9.489 | -104.2214
9.4927 | -104.2152
9.4965 | -104.2065
9.5002 | -104.1946
9.504 | -104.1809
9.5077 | -104.1664
9.5115 | -104.1537
9.5152 | -104.1391
9.519 | -104.1275
9.5227 | -104.116
9.5265 | -104.1017
9.5302 | -104.0858
9.534 | -104.0742
9.5377 | -104.0675
9.5415 | -104.0626
9.5452 | -104.0608
9.549 | -104.0667
9.5527 | -104.0755
9.5565 | -104.0888
9.5602 | -104.1035
9.564 | -104.124
9.5677 | -104.1497
9.5715 | -104.1774
9.5752 | -104.21
9.579 | -104.2424
9.5827 | -104.2709
9.5865 | -104.2965
9.5902 | -104.3236
9.594 | -104.3522
9.5977 | -104.3768
9.6015 | -104.3943
9.6053 | -104.4078
9.609 | -104.417
9.6128 | -104.4191
9.6165 | -104.4115
9.6203 | -104.4032
9.624 | -104.3937
9.6278 | -104.3829
9.6315 | -104.3749
9.6353 | -104.3728
9.639 | -104.3673
9.6428 | -104.3604
9.6465 | -104.3582
9.6503 | -104.3547
9.654 | -104.3549
9.6578 | -104.3591
9.6615 | -104.3659
9.6653 | -104.376
9.669 | -104.3891
9.6728 | -104.3994
9.6765 | -104.4127
9.6803 | -104.4269
9.684 | -104.4396
9.6878 | -104.4522
9.6915 | -104.4613
9.6953 | -104.4657
9.699 | -104.4689
9.7028 | -104.4705
9.7065 | -104.4692
9.7103 | -104.4675
9.714 | -104.4601
9.7178 | -104.4478
9.7215 | -104.4313
9.7253 | -104.4103
9.729 | -104.3905
9.7328 | -104.3677
9.7365 | -104.3345
9.7403 | -104.313
9.744 | -104.2894
9.7478 | -104.2631
9.7515 | -104.2363
9.7553 | -104.2152
9.759 | -104.2002
9.7628 | -104.1873
9.7665 | -104.1784
9.7703 | -104.1722
9.774 | -104.1663
9.7778 | -104.1662
9.7815 | -104.1748
9.7853 | -104.1894
9.789 | -104.2064
9.7928 | -104.2289
9.7965 | -104.2549
9.8003 | -104.2805
9.804 | -104.3106
9.8078 | -104.3449
9.8115 | -104.3804
9.8153 | -104.4117
9.819 | -104.4386
9.8228 | -104.4644
9.8265 | -104.4871
9.8303 | -104.503
9.834 | -104.5191
9.8378 | -104.5326
9.8415 | -104.5391
9.8453 | -104.5383
9.849 | -104.5336
9.8528 | -104.5232
9.8565 | -104.5092
9.8603 | -104.4924
9.864 | -104.4734
9.8678 | -104.4504
9.8715 | -104.4245
9.8753 | -104.3976
9.879 | -104.3664
9.8828 | -104.3303
9.8865 | -104.2904
9.8903 | -104.2502
9.894 | -104.2111
9.8978 | -104.1753
9.9015 | -104.1377
9.9053 | -104.0965
9.909 | -104.0584
9.9128 | -104.0252
9.9165 | -103.994
9.9203 | -103.9689
9.9241 | -103.951
9.9278 | -103.9402
9.9316 | -103.9364
9.9353 | -103.9401
9.9391 | -103.9512
9.9428 | -103.9689
9.9466 | -103.9931
9.9503 | -104.0215
9.9541 | -104.058
9.9578 | -104.094
9.9616 | -104.1331
9.9653 | -104.1748
9.9691 | -104.218
9.9728 | -104.263
9.9766 | -104.3037
9.9803 | -104.3438
9.9841 | -104.3796
9.9878 | -104.4138
9.9916 | -104.4449
9.9953 | -104.471
9.9991 | -104.4946
10.0028 | -104.5175
10.0066 | -104.5365
10.0103 | -104.5518
10.0141 | -104.5663
10.0178 | -104.5765
10.0216 | -104.5849
10.0253 | -104.597
10.0291 | -104.6066
10.0328 | -104.6125
10.0366 | -104.6179
10.0403 | -104.6221
10.0441 | -104.6196
10.0478 | -104.6139
10.0516 | -104.6132
10.0553 | -104.6102
10.0591 | -104.6032
10.0628 | -104.5933
10.0666 | -104.5769
10.0703 | -104.5556
10.0741 | -104.5272
10.0778 | -104.4932
10.0816 | -104.4564
10.0853 | -104.4202
10.0891 | -104.3823
10.0928 | -104.3446
10.0966 | -104.3073
10.1003 | -104.2674
10.1041 | -104.2327
10.1078 | -104.2003
10.1116 | -104.1721
10.1153 | -104.1456
10.1191 | -104.1215
10.1228 | -104.1055
10.1266 | -104.0919
10.1303 | -104.0844
10.1341 | -104.0778
10.1378 | -104.0737
10.1416 | -104.0725
10.1453 | -104.0769
10.1491 | -104.0863
10.1528 | -104.0955
10.1566 | -104.1062
10.1603 | -104.1126
10.1641 | -104.1219
10.1678 | -104.1343
10.1716 | -104.1453
10.1753 | -104.1558
10.1791 | -104.1647
10.1828 | -104.1709
10.1866 | -104.1758
10.1903 | -104.1851
10.1941 | -104.1963
10.1978 | -104.2016
10.2016 | -104.2039
10.2053 | -104.2117
10.2091 | -104.2212
10.2128 | -104.2322
10.2166 | -104.2438
10.2203 | -104.2613
10.2241 | -104.2798
10.2278 | -104.2972
10.2316 | -104.3148
10.2353 | -104.3318
10.2391 | -104.3497
10.2429 | -104.3707
10.2466 | -104.3945
10.2504 | -104.4166
10.2541 | -104.4372
10.2579 | -104.4553
10.2616 | -104.4737
10.2654 | -104.49
10.2691 | -104.5019
10.2729 | -104.5142
10.2766 | -104.5232
10.2804 | -104.5192
10.2841 | -104.524
10.2879 | -104.5269
10.2916 | -104.5275
10.2954 | -104.5239
10.2991 | -104.5185
10.3029 | -104.5154
10.3066 | -104.5109
10.3104 | -104.5087
10.3141 | -104.5038
10.3179 | -104.4996
10.3216 | -104.4925
10.3254 | -104.4841
10.3291 | -104.4801
10.3329 | -104.475
10.3366 | -104.4669
10.3404 | -104.4557
10.3441 | -104.4454
10.3479 | -104.4354
10.3516 | -104.4243
10.3554 | -104.4103
10.3591 | -104.3985
10.3629 | -104.3873
10.3666 | -104.3769
10.3704 | -104.3652
10.3741 | -104.3569
10.3779 | -104.351
10.3816 | -104.3449
10.3854 | -104.3426
10.3891 | -104.3409
10.3929 | -104.3421
10.3966 | -104.3458
10.4004 | -104.3533
10.4041 | -104.3652
10.4079 | -104.3777
10.4116 | -104.3913
10.4154 | -104.4038
10.4191 | -104.4176
10.4229 | -104.434
10.4266 | -104.452
10.4304 | -104.4727
10.4341 | -104.496
10.4379 | -104.5188
10.4416 | -104.5423
10.4454 | -104.5654
10.4491 | -104.5871
10.4529 | -104.607
10.4566 | -104.6276
10.4604 | -104.6488
10.4641 | -104.6682
10.4679 | -104.6822
10.4716 | -104.6937
10.4754 | -104.7069
10.4791 | -104.7199
10.4829 | -104.7248
10.4866 | -104.73
10.4904 | -104.7361
10.4941 | -104.7371
10.4979 | -104.7309
10.5016 | -104.7285
10.5054 | -104.7297
10.5091 | -104.725
10.5129 | -104.7199
10.5166 | -104.7134
10.5204 | -104.7041
10.5241 | -104.6952
10.5279 | -104.6848
10.5316 | -104.6781
10.5354 | -104.6723
10.5391 | -104.6638
10.5429 | -104.6562
10.5466 | -104.6482
10.5504 | -104.6437
10.5541 | -104.6424
10.5579 | -104.6436
10.5617 | -104.6441
10.5654 | -104.6477
10.5692 | -104.6591
10.5729 | -104.6719
10.5767 | -104.6838
10.5804 | -104.6953
10.5842 | -104.7061
10.5879 | -104.7204
10.5917 | -104.7341
10.5954 | -104.7422
10.5992 | -104.7478
10.6029 | -104.7525
10.6067 | -104.7574
10.6104 | -104.76
10.6142 | -104.7602
10.6179 | -104.7594
10.6217 | -104.7585
10.6254 | -104.7572
10.6292 | -104.7536
10.6329 | -104.7503
10.6367 | -104.7458
10.6404 | -104.739
10.6442 | -104.7328
10.6479 | -104.7288
10.6517 | -104.7259
10.6554 | -104.718
10.6592 | -104.7123
10.6629 | -104.7116
10.6667 | -104.7115
10.6704 | -104.7136
10.6742 | -104.7195
10.6779 | -104.726
10.6817 | -104.7375
10.6854 | -104.7479
10.6892 | -104.7601
10.6929 | -104.7727
10.6967 | -104.7837
10.7004 | -104.7953
10.7042 | -104.8063
10.7079 | -104.8186
10.7117 | -104.8291
10.7154 | -104.8405
10.7192 | -104.8479
10.7229 | -104.8495
10.7267 | -104.8503
10.7304 | -104.8507
10.7342 | -104.8562
10.7379 | -104.8638
10.7417 | -104.8684
10.7454 | -104.8766
10.7492 | -104.8909
10.7529 | -104.9045
10.7567 | -104.9159
10.7604 | -104.9224
10.7642 | -104.9317
10.7679 | -104.9404
10.7717 | -104.9486
10.7754 | -104.9599
10.7792 | -104.9673
10.7829 | -104.975
10.7867 | -104.9828
10.7904 | -104.9905
10.7942 | -104.9943
10.7979 | -104.9984
10.8017 | -105.0016
10.8054 | -105.0024
10.8092 | -105.003
10.8129 | -104.9991
10.8167 | -104.9935
10.8204 | -104.9805
10.8242 | -104.9712
10.8279 | -104.9584
10.8317 | -104.9431
10.8354 | -104.9295
10.8392 | -104.9102
10.8429 | -104.8885
10.8467 | -104.8704
10.8504 | -104.8567
10.8542 | -104.8445
10.8579 | -104.8316
10.8617 | -104.8207
10.8654 | -104.8082
10.8692 | -104.7983
10.8729 | -104.7928
10.8767 | -104.7919
10.8804 | -104.7928
10.8842 | -104.7929
10.888 | -104.7934
10.8917 | -104.7925
10.8955 | -104.7941
10.8992 | -104.7954
10.903 | -104.7953
10.9067 | -104.7954
10.9105 | -104.7993
10.9142 | -104.8041
10.918 | -104.8088
10.9217 | -104.8122
10.9255 | -104.8151
10.9292 | -104.8139
10.933 | -104.8076
10.9367 | -104.8025
10.9405 | -104.7976
10.9442 | -104.791
10.948 | -104.7853
10.9517 | -104.7812
10.9555 | -104.7741
10.9592 | -104.7669
10.963 | -104.7597
10.9667 | -104.7506
10.9705 | -104.7424
10.9742 | -104.7366
10.978 | -104.7316
10.9817 | -104.7293
10.9855 | -104.7285
10.9892 | -104.7318
10.993 | -104.7352
10.9967 | -104.7373
11.0005 | -104.7396
11.0042 | -104.743
11.008 | -104.7417
11.0117 | -104.7367
11.0155 | -104.7289
11.0192 | -104.7187
11.023 | -104.707
11.0267 | -104.695
11.0305 | -104.681
11.0342 | -104.6647
11.038 | -104.6454
11.0417 | -104.6241
11.0455 | -104.6026
11.0492 | -104.582
11.053 | -104.561
11.0567 | -104.5413
11.0605 | -104.5288
11.0642 | -104.5198
11.068 | -104.5121
11.0717 | -104.5024
11.0755 | -104.4943
11.0792 | -104.493
11.083 | -104.4903
11.0867 | -104.4906
11.0905 | -104.4921
11.0942 | -104.4938
11.098 | -104.4949
11.1017 | -104.4927
11.1055 | -104.4914
11.1092 | -104.493
11.113 | -104.4958
11.1167 | -104.4947
11.1205 | -104.493
11.1242 | -104.4964
11.128 | -104.4975
11.1317 | -104.4973
11.1355 | -104.501
11.1392 | -104.5053
11.143 | -104.5104
11.1467 | -104.5151
11.1505 | -104.5221
11.1542 | -104.5371
11.158 | -104.5531
11.1617 | -104.5692
11.1655 | -104.5865
11.1692 | -104.6073
11.173 | -104.6308
11.1767 | -104.6559
11.1805 | -104.6813
11.1842 | -104.7052
11.188 | -104.7283
11.1917 | -104.7512
11.1955 | -104.7728
11.1992 | -104.796
11.203 | -104.8195
11.2068 | -104.8382
11.2105 | -104.8526
11.2143 | -104.868
11.218 | -104.8787
11.2218 | -104.8841
11.2255 | -104.8903
11.2293 | -104.8973
11.233 | -104.9061
11.2368 | -104.9138
11.2405 | -104.9168
11.2443 | -104.9211
11.248 | -104.9285
11.2518 | -104.9401
11.2555 | -104.9521
11.2593 | -104.9655
11.263 | -104.9845
11.2668 | -105.0016
11.2705 | -105.0166
11.2743 | -105.0342
11.278 | -105.0563
11.2818 | -105.082
11.2855 | -105.1065
11.2893 | -105.1371
11.293 | -105.1667
11.2968 | -105.1932
11.3005 | -105.2259
11.3043 | -105.2568
11.308 | -105.2884
11.3118 | -105.3133
11.3155 | -105.3321
11.3193 | -105.3489
11.323 | -105.3702
11.3268 | -105.3865
11.3305 | -105.3962
11.3343 | -105.4016
11.338 | -105.405
11.3418 | -105.4035
11.3455 | -105.4001
11.3493 | -105.392
11.353 | -105.3807
11.3568 | -105.3681
11.3605 | -105.3452
11.3643 | -105.3234
11.368 | -105.2982
11.3718 | -105.2751
11.3755 | -105.2489
11.3793 | -105.22
11.383 | -105.1894
11.3868 | -105.1607
11.3905 | -105.1342
11.3943 | -105.106
11.398 | -105.0774
11.4018 | -105.0521
11.4055 | -105.0315
11.4093 | -105.0112
11.413 | -104.9921
11.4168 | -104.9775
11.4205 | -104.9673
11.4243 | -104.9605
11.428 | -104.957
11.4318 | -104.957
11.4355 | -104.9596
11.4393 | -104.9659
11.443 | -104.9754
11.4468 | -104.9837
11.4505 | -104.9934
11.4543 | -105.0059
11.458 | -105.0202
11.4618 | -105.0323
11.4655 | -105.0414
11.4693 | -105.0523
11.473 | -105.0617
11.4768 | -105.0715
11.4805 | -105.0765
11.4843 | -105.0818
11.488 | -105.0823
11.4918 | -105.0818
11.4955 | -105.081
11.4993 | -105.0783
11.503 | -105.0715
11.5068 | -105.0625
11.5105 | -105.0521
11.5143 | -105.0418
11.5181 | -105.0296
11.5218 | -105.018
11.5256 | -105.0065
11.5293 | -104.9971
11.5331 | -104.994
11.5368 | -104.9926
11.5406 | -104.988
11.5443 | -104.983
11.5481 | -104.9802
11.5518 | -104.9821
11.5556 | -104.9835
11.5593 | -104.9822
11.5631 | -104.9849
11.5668 | -104.9904
11.5706 | -104.9948
11.5743 | -104.9927
11.5781 | -104.9882
11.5818 | -104.9828
11.5856 | -104.9774
11.5893 | -104.971
11.5931 | -104.9636
11.5968 | -104.9546
11.6006 | -104.9437
11.6043 | -104.9331
11.6081 | -104.9254
11.6118 | -104.9187
11.6156 | -104.9125
11.6193 | -104.9059
11.6231 | -104.9006
11.6268 | -104.8992
11.6306 | -104.8981
11.6343 | -104.8985
11.6381 | -104.9015
11.6418 | -104.9052
11.6456 | -104.9073
11.6493 | -104.9191
11.6531 | -104.9329
11.6568 | -104.9451
11.6606 | -104.9605
11.6643 | -104.9773
11.6681 | -104.9962
11.6718 | -105.0147
11.6756 | -105.0336
11.6793 | -105.051
11.6831 | -105.0694
11.6868 | -105.0868
11.6906 | -105.0977
11.6943 | -105.1079
11.6981 | -105.1179
11.7018 | -105.1233
11.7056 | -105.1249
11.7093 | -105.1287
11.7131 | -105.1331
11.7168 | -105.1318
11.7206 | -105.125
11.7243 | -105.1177
11.7281 | -105.1122
11.7318 | -105.1047
11.7356 | -105.092
11.7393 | -105.0772
11.7431 | -105.0649
11.7468 | -105.0551
11.7506 | -105.0412
11.7543 | -105.0243
11.7581 | -105.0071
11.7618 | -104.9871
11.7656 | -104.9652
11.7693 | -104.9463
11.7731 | -104.927
11.7768 | -104.9028
11.7806 | -104.8822
11.7843 | -104.8626
11.7881 | -104.8431
11.7918 | -104.824
11.7956 | -104.8071
11.7993 | -104.793
11.8031 | -104.7793
11.8068 | -104.7661
11.8106 | -104.756
11.8143 | -104.75
11.8181 | -104.7376
11.8218 | -104.7279
11.8256 | -104.726
11.8293 | -104.7242
11.8331 | -104.725
11.8369 | -104.7279
11.8406 | -104.7345
11.8444 | -104.7463
11.8481 | -104.7627
11.8519 | -104.7805
11.8556 | -104.8015
11.8594 | -104.8291
11.8631 | -104.8584
11.8669 | -104.8895
11.8706 | -104.9216
11.8744 | -104.9561
11.8781 | -104.986
11.8819 | -105.0145
11.8856 | -105.0445
11.8894 | -105.0741
11.8931 | -105.1022
11.8969 | -105.1239
11.9006 | -105.1315
11.9044 | -105.1499
11.9081 | -105.1671
11.9119 | -105.1808
11.9156 | -105.1918
11.9194 | -105.1951
11.9231 | -105.1981
11.9269 | -105.2
11.9306 | -105.1994
11.9344 | -105.1972
11.9381 | -105.1928
11.9419 | -105.189
11.9456 | -105.1843
11.9494 | -105.1785
11.9531 | -105.1687
11.9569 | -105.1639
11.9606 | -105.16
11.9644 | -105.1552
11.9681 | -105.1505
11.9719 | -105.1486
11.9756 | -105.15
11.9794 | -105.1512
11.9831 | -105.1513
11.9869 | -105.1555
11.9906 | -105.1635
11.9944 | -105.1724
11.9981 | -105.1781
12.0019 | -105.1849
12.0056 | -105.1917
12.0094 | -105.1964
12.0131 | -105.2037
12.0169 | -105.211
12.0206 | -105.2159
12.0244 | -105.2191
12.0281 | -105.2209
12.0319 | -105.2221
12.0356 | -105.2198
12.0394 | -105.2162
12.0431 | -105.2167
12.0469 | -105.2176
12.0506 | -105.2196
12.0544 | -105.2209
12.0581 | -105.2186
12.0619 | -105.2164
12.0656 | -105.2172
12.0694 | -105.2172
12.0731 | -105.2183
12.0769 | -105.2178
12.0806 | -105.2177
12.0844 | -105.2197
12.0881 | -105.2208
12.0919 | -105.2208
12.0956 | -105.2205
12.0994 | -105.2189
12.1031 | -105.2159
12.1069 | -105.2155
12.1106 | -105.2187
12.1144 | -105.2198
12.1181 | -105.2189
12.1219 | -105.2189
12.1256 | -105.218
12.1294 | -105.2191
12.1331 | -105.2193
12.1369 | -105.223
12.1406 | -105.2293
12.1444 | -105.2348
12.1481 | -105.2404
12.1519 | -105.2452
12.1556 | -105.2559
12.1594 | -105.2692
12.1631 | -105.281
12.1669 | -105.2945
12.1707 | -105.3102
12.1744 | -105.3281
12.1782 | -105.3478
12.1819 | -105.368
12.1857 | -105.3889
12.1894 | -105.4071
12.1932 | -105.4203
12.1969 | -105.4311
12.2007 | -105.4441
12.2044 | -105.4577
12.2082 | -105.4682
12.2119 | -105.4771
12.2157 | -105.4843
12.2194 | -105.4859
12.2232 | -105.484
12.2269 | -105.4799
12.2307 | -105.4752
12.2344 | -105.466
12.2382 | -105.4531
12.2419 | -105.4375
12.2457 | -105.4226
12.2494 | -105.4017
12.2532 | -105.3791
12.2569 | -105.357
12.2607 | -105.3326
12.2644 | -105.3074
12.2682 | -105.2807
12.2719 | -105.2562
12.2757 | -105.2309
12.2794 | -105.2068
12.2832 | -105.1812
12.2869 | -105.1567
12.2907 | -105.1405
12.2944 | -105.1257
12.2982 | -105.1131
12.3019 | -105.1058
12.3057 | -105.1009
12.3094 | -105.0977
12.3132 | -105.0946
12.3169 | -105.0945
12.3207 | -105.1013
12.3244 | -105.1111
12.3282 | -105.1192
12.3319 | -105.1281
12.3357 | -105.1411
12.3394 | -105.1558
12.3432 | -105.169
12.3469 | -105.1811
12.3507 | -105.1983
12.3544 | -105.2108
12.3582 | -105.2213
12.3619 | -105.232
12.3657 | -105.2432
12.3694 | -105.256
12.3732 | -105.2663
12.3769 | -105.273
12.3807 | -105.2779
12.3844 | -105.2827
12.3882 | -105.2889
12.3919 | -105.2929
12.3957 | -105.2958
12.3994 | -105.3
12.4032 | -105.3043
12.4069 | -105.3097
12.4107 | -105.3196
12.4144 | -105.3266
12.4182 | -105.3311
12.4219 | -105.3356
12.4257 | -105.3444
12.4294 | -105.3527
12.4332 | -105.3593
12.4369 | -105.3695
12.4407 | -105.3825
12.4444 | -105.3977
12.4482 | -105.4072
12.4519 | -105.4153
12.4557 | -105.423
12.4594 | -105.4307
12.4632 | -105.4371
12.4669 | -105.4401
12.4707 | -105.4411
12.4744 | -105.4411
12.4782 | -105.4419
12.4819 | -105.4419
12.4857 | -105.44
12.4895 | -105.4366
12.4932 | -105.436
12.497 | -105.4353
12.5007 | -105.4293
12.5045 | -105.4244
12.5082 | -105.424
12.512 | -105.4227
12.5157 | -105.4194
12.5195 | -105.4163
12.5232 | -105.4149
12.527 | -105.4139
12.5307 | -105.4123
12.5345 | -105.4164
12.5382 | -105.4243
12.542 | -105.4254
12.5457 | -105.4235
12.5495 | -105.4254
12.5532 | -105.4332
12.557 | -105.4401
12.5607 | -105.4456
12.5645 | -105.4519
12.5682 | -105.4566
12.572 | -105.4596
12.5757 | -105.4606
12.5795 | -105.4562
12.5832 | -105.4529
12.587 | -105.4518
12.5907 | -105.4479
12.5945 | -105.4446
12.5982 | -105.4388
12.602 | -105.4297
12.6057 | -105.4181
12.6095 | -105.4027
12.6132 | -105.3848
12.617 | -105.363
12.6207 | -105.338
12.6245 | -105.3106
12.6282 | -105.2844
12.632 | -105.2661
12.6357 | -105.2476
12.6395 | -105.2291
12.6432 | -105.2091
12.647 | -105.1907
12.6507 | -105.1758
12.6545 | -105.1646
12.6582 | -105.159
12.662 | -105.153
12.6657 | -105.1482
12.6695 | -105.1464
12.6732 | -105.1481
12.677 | -105.1526
12.6807 | -105.1573
12.6845 | -105.1672
12.6882 | -105.1787
12.692 | -105.1915
12.6957 | -105.201
12.6995 | -105.2102
12.7032 | -105.2212
12.707 | -105.2346
12.7107 | -105.2464
12.7145 | -105.258
12.7182 | -105.2713
12.722 | -105.283
12.7257 | -105.2964
12.7295 | -105.3054
12.7332 | -105.3092
12.737 | -105.309
12.7407 | -105.3086
12.7445 | -105.3092
12.7482 | -105.3076
12.752 | -105.305
12.7557 | -105.3017
12.7595 | -105.2944
12.7632 | -105.2898
12.767 | -105.2807
12.7707 | -105.2639
12.7745 | -105.2466
12.7782 | -105.2277
12.782 | -105.2103
12.7857 | -105.1966
12.7895 | -105.1802
12.7932 | -105.1666
12.797 | -105.1538
12.8008 | -105.1424
12.8045 | -105.131
12.8083 | -105.1206
12.812 | -105.106
12.8158 | -105.0896
12.8195 | -105.0783
12.8233 | -105.0699
12.827 | -105.0611
12.8308 | -105.0521
12.8345 | -105.0434
12.8383 | -105.0335
12.842 | -105.0233
12.8458 | -105.0176
12.8495 | -105.0148
12.8533 | -105.013
12.857 | -105.0135
12.8608 | -105.012
12.8645 | -105.0121
12.8683 | -105.0121
12.872 | -105.0126
12.8758 | -105.0123
12.8795 | -105.0195
12.8833 | -105.029
12.887 | -105.037
12.8908 | -105.0408
12.8945 | -105.0482
12.8983 | -105.0532
12.902 | -105.056
12.9058 | -105.0579
12.9095 | -105.0633
12.9133 | -105.0692
12.917 | -105.0724
12.9208 | -105.0758
12.9245 | -105.078
12.9283 | -105.0784
12.932 | -105.0781
12.9358 | -105.0758
12.9395 | -105.0715
12.9433 | -105.0698
12.947 | -105.0671
12.9508 | -105.0635
12.9545 | -105.0624
12.9583 | -105.0627
12.962 | -105.0635
12.9658 | -105.0648
12.9695 | -105.0657
12.9733 | -105.0684
12.977 | -105.0709
12.9808 | -105.0742
12.9845 | -105.0851
12.9883 | -105.0927
12.992 | -105.1005
12.9958 | -105.1075
12.9995 | -105.1195
13.0033 | -105.1304
13.007 | -105.1373
13.0108 | -105.1449
13.0145 | -105.1559
13.0183 | -105.1684
13.022 | -105.1775
13.0258 | -105.1868
13.0295 | -105.194
13.0333 | -105.1997
13.037 | -105.2058
13.0408 | -105.2108
13.0445 | -105.2134
13.0483 | -105.2212
13.052 | -105.2277
13.0558 | -105.2321
13.0595 | -105.236
13.0633 | -105.2386
13.067 | -105.2378
13.0708 | -105.2385
13.0745 | -105.2396
13.0783 | -105.2441
13.082 | -105.2462
13.0858 | -105.2483
13.0895 | -105.2524
13.0933 | -105.2568
13.097 | -105.2593
13.1008 | -105.2609
13.1045 | -105.2633
13.1083 | -105.2652
13.112 | -105.2676
13.1158 | -105.2693
13.1196 | -105.2712
13.1233 | -105.2766
13.1271 | -105.2855
13.1308 | -105.2949
13.1346 | -105.2981
13.1383 | -105.3027
13.1421 | -105.3112
13.1458 | -105.3238
13.1496 | -105.3372
13.1533 | -105.3487
13.1571 | -105.3636
13.1608 | -105.3805
13.1646 | -105.3989
13.1683 | -105.4192
13.1721 | -105.4406
13.1758 | -105.4594
13.1796 | -105.4792
13.1833 | -105.5009
13.1871 | -105.5238
13.1908 | -105.5423
13.1946 | -105.5624
13.1983 | -105.5809
13.2021 | -105.5975
13.2058 | -105.61
13.2096 | -105.6219
13.2133 | -105.6349
13.2171 | -105.6459
13.2208 | -105.6543
13.2246 | -105.6607
13.2283 | -105.6615
13.2321 | -105.6567
13.2358 | -105.6485
13.2396 | -105.6351
13.2433 | -105.623
13.2471 | -105.6111
13.2508 | -105.5931
13.2546 | -105.5737
13.2583 | -105.5522
13.2621 | -105.5276
13.2658 | -105.5014
13.2696 | -105.4739
13.2733 | -105.4446
13.2771 | -105.4134
13.2808 | -105.3818
13.2846 | -105.353
13.2883 | -105.3242
13.2921 | -105.2953
13.2958 | -105.2671
13.2996 | -105.2407
13.3033 | -105.2156
13.3071 | -105.1915
13.3108 | -105.1701
13.3146 | -105.1484
13.3183 | -105.1255
13.3221 | -105.1041
13.3258 | -105.092
13.3296 | -105.08
13.3333 | -105.0665
13.3371 | -105.0518
13.3408 | -105.0387
13.3446 | -105.0285
13.3483 | -105.02
13.3521 | -105.0119
13.3558 | -105.0045
13.3596 | -104.9998
13.3633 | -104.997
13.3671 | -104.9956
13.3708 | -104.9944
13.3746 | -104.9963
13.3783 | -104.9984
13.3821 | -105.0021
13.3858 | -105.0048
13.3896 | -105.0106
13.3933 | -105.0182
13.3971 | -105.0278
13.4008 | -105.034
13.4046 | -105.0421
13.4083 | -105.0505
13.4121 | -105.0623
13.4158 | -105.0736
13.4196 | -105.0873
13.4233 | -105.1055
13.4271 | -105.1252
13.4308 | -105.1426
13.4346 | -105.1586
13.4383 | -105.1758
13.4421 | -105.1975
13.4459 | -105.2203
13.4496 | -105.2433
13.4534 | -105.267
13.4571 | -105.2916
13.4609 | -105.3174
13.4646 | -105.3447
13.4684 | -105.3745
13.4721 | -105.4025
13.4759 | -105.4296
13.4796 | -105.4586
13.4834 | -105.4873
13.4871 | -105.5128
13.4909 | -105.5348
13.4946 | -105.559
13.4984 | -105.5849
13.5021 | -105.6087
13.5059 | -105.6311
13.5096 | -105.6477
13.5134 | -105.661
13.5171 | -105.6717
13.5209 | -105.6835
13.5246 | -105.6795
13.5284 | -105.6861
13.5321 | -105.69
13.5359 | -105.6919
13.5396 | -105.69
13.5434 | -105.6855
13.5471 | -105.6817
13.5509 | -105.6747
13.5546 | -105.6612
13.5584 | -105.6459
13.5621 | -105.6307
13.5659 | -105.6156
13.5696 | -105.5971
13.5734 | -105.58
13.5771 | -105.5636
13.5809 | -105.5463
13.5846 | -105.5285
13.5884 | -105.5116
13.5921 | -105.4935
13.5959 | -105.4792
13.5996 | -105.4622
13.6034 | -105.4469
13.6071 | -105.433
13.6109 | -105.4212
13.6146 | -105.4129
13.6184 | -105.4029
13.6221 | -105.3915
13.6259 | -105.3819
13.6296 | -105.372
13.6334 | -105.3635
13.6371 | -105.3525
13.6409 | -105.3402
13.6446 | -105.3309
13.6484 | -105.3226
13.6521 | -105.3153
13.6559 | -105.3063
13.6596 | -105.2984
13.6634 | -105.2908
13.6671 | -105.2834
13.6709 | -105.2768
13.6746 | -105.2696
13.6784 | -105.2662
13.6821 | -105.2633
13.6859 | -105.2595
13.6896 | -105.2568
13.6934 | -105.2567
13.6971 | -105.2585
13.7009 | -105.2603
13.7046 | -105.2669
13.7084 | -105.2772
13.7121 | -105.2848
13.7159 | -105.2896
13.7196 | -105.297
13.7234 | -105.3068
13.7271 | -105.3133
13.7309 | -105.3209
13.7346 | -105.3289
13.7384 | -105.3368
13.7421 | -105.3426
13.7459 | -105.348
13.7496 | -105.3562
13.7534 | -105.3655
13.7571 | -105.3721
13.7609 | -105.3762
13.7647 | -105.3788
13.7684 | -105.3833
13.7722 | -105.3883
13.7759 | -105.3936
13.7797 | -105.4001
13.7834 | -105.4059
13.7872 | -105.4113
13.7909 | -105.4203
13.7947 | -105.4303
13.7984 | -105.4416
13.8022 | -105.4521
13.8059 | -105.4625
13.8097 | -105.4728
13.8134 | -105.4818
13.8172 | -105.4909
13.8209 | -105.4977
13.8247 | -105.5042
13.8284 | -105.5148
13.8322 | -105.5254
13.8359 | -105.5354
13.8397 | -105.5446
13.8434 | -105.5477
13.8472 | -105.5488
13.8509 | -105.5501
13.8547 | -105.5495
13.8584 | -105.5462
13.8622 | -105.5406
13.8659 | -105.5322
13.8697 | -105.5235
13.8734 | -105.5148
13.8772 | -105.5014
13.8809 | -105.4878
13.8847 | -105.4758
13.8884 | -105.4634
13.8922 | -105.45
13.8959 | -105.4338
13.8997 | -105.4206
13.9034 | -105.4113
13.9072 | -105.3984
13.9109 | -105.3882
13.9147 | -105.382
13.9184 | -105.3737
13.9222 | -105.3664
13.9259 | -105.3606
13.9297 | -105.3552
13.9334 | -105.3514
13.9372 | -105.3452
13.9409 | -105.3428
13.9447 | -105.341
13.9484 | -105.3361
13.9522 | -105.3348
13.9559 | -105.3352
13.9597 | -105.3342
13.9634 | -105.3329
13.9672 | -105.3343
13.9709 | -105.3335
13.9747 | -105.3321
13.9784 | -105.3298
13.9822 | -105.33
13.9859 | -105.331
13.9897 | -105.3287
13.9934 | -105.3252
13.9972 | -105.3237
14.0009 | -105.3228
14.0047 | -105.3227
14.0084 | -105.3225
14.0122 | -105.3218
14.0159 | -105.3246
14.0197 | -105.3315
14.0234 | -105.3358
14.0272 | -105.337
14.0309 | -105.3358
14.0347 | -105.3373
14.0384 | -105.3417
14.0422 | -105.3487
14.0459 | -105.3563
14.0497 | -105.3602
14.0534 | -105.3683
14.0572 | -105.3756
14.0609 | -105.3816
14.0647 | -105.386
14.0684 | -105.3945
14.0722 | -105.4007
14.0759 | -105.4043
14.0797 | -105.4082
14.0835 | -105.409
14.0872 | -105.4088
14.091 | -105.4059
14.0947 | -105.4022
14.0985 | -105.4012
14.1022 | -105.3989
14.106 | -105.3931
14.1097 | -105.3881
14.1135 | -105.3815
14.1172 | -105.3714
14.121 | -105.365
14.1247 | -105.3586
14.1285 | -105.3522
14.1322 | -105.3437
14.136 | -105.3357
14.1397 | -105.3324
14.1435 | -105.3271
14.1472 | -105.3209
14.151 | -105.3171
14.1547 | -105.3139
14.1585 | -105.3115
14.1622 | -105.3131
14.166 | -105.3148
14.1697 | -105.3144
14.1735 | -105.3161
14.1772 | -105.3191
14.181 | -105.3268
14.1847 | -105.336
14.1885 | -105.3455
14.1922 | -105.3563
14.196 | -105.3676
14.1997 | -105.3798
14.2035 | -105.3939
14.2072 | -105.4091
14.211 | -105.4243
14.2147 | -105.4398
14.2185 | -105.4542
14.2222 | -105.4669
14.226 | -105.4829
14.2297 | -105.4989
14.2335 | -105.5138
14.2372 | -105.5269
14.241 | -105.541
14.2447 | -105.5559
14.2485 | -105.5704
14.2522 | -105.5874
14.256 | -105.6006
14.2597 | -105.6132
14.2635 | -105.6264
14.2672 | -105.6378
14.271 | -105.6479
14.2747 | -105.6582
14.2785 | -105.6693
14.2822 | -105.6805
14.286 | -105.6912
14.2897 | -105.7053
14.2935 | -105.7154
14.2972 | -105.7252
14.301 | -105.7374
14.3047 | -105.7472
14.3085 | -105.7584
14.3122 | -105.7717
14.316 | -105.7855
14.3197 | -105.798
14.3235 | -105.8108
14.3272 | -105.8198
14.331 | -105.8309
14.3347 | -105.8403
14.3385 | -105.8523
14.3422 | -105.8605
14.346 | -105.8643
14.3497 | -105.8686
14.3535 | -105.8747
14.3572 | -105.8805
14.361 | -105.8809
14.3647 | -105.8812
14.3685 | -105.883
14.3722 | -105.8838
14.376 | -105.8814
14.3797 | -105.8734
14.3835 | -105.8666
14.3872 | -105.8558
14.391 | -105.8414
14.3947 | -105.825
14.3985 | -105.807
14.4023 | -105.7858
14.406 | -105.7644
14.4098 | -105.7441
14.4135 | -105.724
14.4173 | -105.7036
14.421 | -105.6805
14.4248 | -105.6582
14.4285 | -105.6367
14.4323 | -105.613
14.436 | -105.5906
14.4398 | -105.5707
14.4435 | -105.5496
14.4473 | -105.5268
14.451 | -105.5056
14.4548 | -105.4874
14.4585 | -105.4731
14.4623 | -105.4592
14.466 | -105.4484
14.4698 | -105.4391
14.4735 | -105.4309
14.4773 | -105.4252
14.481 | -105.4182
14.4848 | -105.4132
14.4885 | -105.4113
14.4923 | -105.4102
14.496 | -105.4091
14.4998 | -105.4102
14.5035 | -105.4131
14.5073 | -105.4134
14.511 | -105.4148
14.5148 | -105.4168
14.5185 | -105.4188
14.5223 | -105.4232
14.526 | -105.4311
14.5298 | -105.4418
14.5335 | -105.4526
14.5373 | -105.4614
14.541 | -105.4702
14.5448 | -105.478
14.5485 | -105.4887
14.5523 | -105.4991
14.556 | -105.5103
14.5598 | -105.5196
14.5635 | -105.5304
14.5673 | -105.5456
14.571 | -105.5618
14.5748 | -105.5756
14.5785 | -105.5913
14.5823 | -105.6084
14.586 | -105.6297
14.5898 | -105.6492
14.5935 | -105.6695
14.5973 | -105.6903
14.601 | -105.7121
14.6048 | -105.7345
14.6085 | -105.7558
14.6123 | -105.7765
14.616 | -105.7934
14.6198 | -105.8103
14.6235 | -105.8264
14.6273 | -105.842
14.631 | -105.8582
14.6348 | -105.8724
14.6385 | -105.8872
14.6423 | -105.9006
14.646 | -105.9091
14.6498 | -105.9172
14.6535 | -105.9272
14.6573 | -105.9369
14.661 | -105.9429
14.6648 | -105.9452
14.6685 | -105.9434
14.6723 | -105.9413
14.676 | -105.9376
14.6798 | -105.9297
14.6835 | -105.9215
14.6873 | -105.9124
14.691 | -105.9027
14.6948 | -105.8886
14.6985 | -105.877
14.7023 | -105.8664
14.706 | -105.856
14.7098 | -105.8441
14.7135 | -105.8282
14.7173 | -105.8124
14.721 | -105.7992
14.7248 | -105.7838
14.7286 | -105.766
14.7323 | -105.7486
14.7361 | -105.7297
14.7398 | -105.7131
14.7436 | -105.6952
14.7473 | -105.6776
14.7511 | -105.6638
14.7548 | -105.6521
14.7586 | -105.6394
14.7623 | -105.6285
14.7661 | -105.6149
14.7698 | -105.6034
14.7736 | -105.5903
14.7773 | -105.5759
14.7811 | -105.5606
14.7848 | -105.5495
14.7886 | -105.5375
14.7923 | -105.525
14.7961 | -105.5149
14.7998 | -105.5044
14.8036 | -105.493
14.8073 | -105.4814
14.8111 | -105.4707
14.8148 | -105.4612
14.8186 | -105.4519
14.8223 | -105.4418
14.8261 | -105.4333
14.8298 | -105.4262
14.8336 | -105.4158
14.8373 | -105.4082
14.8411 | -105.4023
14.8448 | -105.3963
14.8486 | -105.3875
14.8523 | -105.3774
14.8561 | -105.3694
14.8598 | -105.3635
14.8636 | -105.3566
14.8673 | -105.3476
14.8711 | -105.3377
14.8748 | -105.326
14.8786 | -105.3136
14.8823 | -105.3043
14.8861 | -105.2943
14.8898 | -105.284
14.8936 | -105.2754
14.8973 | -105.2628
14.9011 | -105.2497
14.9048 | -105.2352
14.9086 | -105.222
14.9123 | -105.2045
14.9161 | -105.1887
14.9198 | -105.1732
14.9236 | -105.1589
14.9273 | -105.1459
14.9311 | -105.1336
14.9348 | -105.1218
14.9386 | -105.1107
14.9423 | -105.1014
14.9461 | -105.0902
14.9498 | -105.0804
14.9536 | -105.0727
14.9573 | -105.0696
14.9611 | -105.0672
14.9648 | -105.065
14.9686 | -105.0636
14.9723 | -105.0644
14.9761 | -105.0663
14.9798 | -105.0709
14.9836 | -105.077
14.9873 | -105.0861
14.9911 | -105.0967
14.9948 | -105.1091
14.9986 | -105.122
15.0023 | -105.1378
15.0061 | -105.1561
15.0098 | -105.1732
15.0136 | -105.193
15.0173 | -105.2134
15.0211 | -105.2357
15.0248 | -105.2581
15.0286 | -105.2815
15.0323 | -105.304
15.0361 | -105.328
15.0398 | -105.3514
15.0436 | -105.3734
15.0474 | -105.3958
15.0511 | -105.4155
15.0549 | -105.4365
15.0586 | -105.4564
15.0624 | -105.4754
15.0661 | -105.4938
15.0699 | -105.5094
15.0736 | -105.5228
15.0774 | -105.5345
15.0811 | -105.544
15.0849 | -105.5514
15.0886 | -105.5592
15.0924 | -105.5645
15.0961 | -105.5671
15.0999 | -105.5684
15.1036 | -105.5662
15.1074 | -105.5655
15.1111 | -105.5669
15.1149 | -105.5621
15.1186 | -105.556
15.1224 | -105.5469
15.1261 | -105.5365
15.1299 | -105.5259
15.1336 | -105.5126
15.1374 | -105.4975
15.1411 | -105.4828
15.1449 | -105.4662
15.1486 | -105.4412
15.1524 | -105.4236
15.1561 | -105.4047
15.1599 | -105.3855
15.1636 | -105.3672
15.1674 | -105.348
15.1711 | -105.3279
15.1749 | -105.3093
15.1786 | -105.2916
15.1824 | -105.2736
15.1861 | -105.2582
15.1899 | -105.244
15.1936 | -105.2317
15.1974 | -105.2207
15.2011 | -105.2112
15.2049 | -105.2056
15.2086 | -105.2007
15.2124 | -105.1955
15.2161 | -105.1869
15.2199 | -105.1839
15.2236 | -105.1844
15.2274 | -105.1855
15.2311 | -105.1867
15.2349 | -105.1905
15.2386 | -105.1961
15.2424 | -105.203
15.2461 | -105.2102
15.2499 | -105.2173
15.2536 | -105.2253
15.2574 | -105.2335
15.2611 | -105.2389
15.2649 | -105.2435
15.2686 | -105.2495
15.2724 | -105.2548
15.2761 | -105.2586
15.2799 | -105.2631
15.2836 | -105.265
15.2874 | -105.2634
15.2911 | -105.2598
15.2949 | -105.2564
15.2986 | -105.2528
15.3024 | -105.2477
15.3061 | -105.2424
15.3099 | -105.2374
15.3136 | -105.2284
15.3174 | -105.2192
15.3211 | -105.2086
15.3249 | -105.1962
15.3286 | -105.1845
15.3324 | -105.1767
15.3361 | -105.1685
15.3399 | -105.161
15.3436 | -105.1549
15.3474 | -105.1448
15.3511 | -105.135
15.3549 | -105.1287
15.3587 | -105.1233
15.3624 | -105.1182
15.3662 | -105.1145
15.3699 | -105.1099
15.3737 | -105.1041
15.3774 | -105.0996
15.3812 | -105.0978
15.3849 | -105.0978
15.3887 | -105.0983
15.3924 | -105.0976
15.3962 | -105.0988
15.3999 | -105.1024
15.4037 | -105.1064
15.4074 | -105.1088
15.4112 | -105.1121
15.4149 | -105.118
15.4187 | -105.1252
15.4224 | -105.1342
15.4262 | -105.1419
15.4299 | -105.1501
15.4337 | -105.162
15.4374 | -105.1779
15.4412 | -105.1877
15.4449 | -105.1953
15.4487 | -105.2036
15.4524 | -105.2128
15.4562 | -105.2233
15.4599 | -105.2321
15.4637 | -105.2407
15.4674 | -105.25
15.4712 | -105.2575
15.4749 | -105.265
15.4787 | -105.2739
15.4824 | -105.2836
15.4862 | -105.2943
15.4899 | -105.3034
15.4937 | -105.3094
15.4974 | -105.3162
15.5012 | -105.3235
15.5049 | -105.3301
15.5087 | -105.3378
15.5124 | -105.3441
15.5162 | -105.3498
15.5199 | -105.3553
15.5237 | -105.359
15.5274 | -105.3611
15.5312 | -105.3624
15.5349 | -105.3623
15.5387 | -105.3622
15.5424 | -105.3624
15.5462 | -105.3614
15.5499 | -105.3595
15.5537 | -105.3567
15.5574 | -105.3514
15.5612 | -105.3449
15.5649 | -105.3361
15.5687 | -105.328
15.5724 | -105.3225
15.5762 | -105.3139
15.5799 | -105.3057
15.5837 | -105.299
15.5874 | -105.2907
15.5912 | -105.2813
15.5949 | -105.275
15.5987 | -105.2703
15.6024 | -105.2621
15.6062 | -105.2552
15.6099 | -105.2478
15.6137 | -105.2388
15.6174 | -105.2308
15.6212 | -105.2235
15.6249 | -105.217
15.6287 | -105.2108
15.6324 | -105.2067
15.6362 | -105.2028
15.6399 | -105.1986
15.6437 | -105.1944
15.6474 | -105.1906
15.6512 | -105.1847
15.6549 | -105.1747
15.6587 | -105.1665
15.6624 | -105.1605
15.6662 | -105.154
15.6699 | -105.1455
15.6737 | -105.1364
15.6775 | -105.1271
15.6812 | -105.1171
15.685 | -105.1069
15.6887 | -105.096
15.6925 | -105.0825
15.6962 | -105.0691
15.7 | -105.0574
15.7037 | -105.0466
15.7075 | -105.0353
15.7112 | -105.0224
15.715 | -105.0089
15.7187 | -104.9964
15.7225 | -104.9855
15.7262 | -104.9746
15.73 | -104.9645
15.7337 | -104.9541
15.7375 | -104.9476
15.7412 | -104.9392
15.745 | -104.9297
15.7487 | -104.9186
15.7525 | -104.9093
15.7562 | -104.9023
15.76 | -104.8992
15.7637 | -104.8965
15.7675 | -104.8953
15.7712 | -104.8933
15.775 | -104.8923
15.7787 | -104.8913
15.7825 | -104.8915
15.7862 | -104.8919
15.79 | -104.8948
15.7937 | -104.8975
15.7975 | -104.8974
15.8012 | -104.8961
15.805 | -104.894
15.8087 | -104.8945
15.8125 | -104.8947
15.8162 | -104.8952
15.82 | -104.8911
15.8237 | -104.8848
15.8275 | -104.883
15.8312 | -104.8777
15.835 | -104.8728
15.8387 | -104.8689
15.8425 | -104.862
15.8462 | -104.8545
15.85 | -104.8471
15.8537 | -104.8384
15.8575 | -104.8245
15.8612 | -104.8098
15.865 | -104.7958
15.8687 | -104.783
15.8725 | -104.7707
15.8762 | -104.7633
15.88 | -104.7605
15.8837 | -104.7543
15.8875 | -104.7472
15.8912 | -104.7421
15.895 | -104.7403
15.8987 | -104.7373
15.9025 | -104.7344
15.9062 | -104.7344
15.91 | -104.7353
15.9137 | -104.7373
15.9175 | -104.7439
15.9212 | -104.7541
15.925 | -104.7642
15.9287 | -104.7726
15.9325 | -104.7804
15.9362 | -104.791
15.94 | -104.8034
15.9437 | -104.8146
15.9475 | -104.8298
15.9512 | -104.8437
15.955 | -104.8554
15.9587 | -104.8654
15.9625 | -104.8763
15.9662 | -104.8873
15.97 | -104.8986
15.9737 | -104.9098
15.9775 | -104.9222
15.9812 | -104.9302
15.985 | -104.9363
15.9887 | -104.9389
15.9925 | -104.9408
15.9962 | -104.9416
16 | -104.943
16.0038 | -104.9412
16.0075 | -104.9405
16.0113 | -104.9411
16.015 | -104.9381
16.0188 | -104.9316
16.0225 | -104.9264
16.0263 | -104.9221
16.03 | -104.922
16.0338 | -104.919
16.0375 | -104.9128
16.0413 | -104.9088
16.045 | -104.9062
16.0488 | -104.9054
16.0525 | -104.9035
16.0563 | -104.9009
16.06 | -104.9013
16.0638 | -104.9027
16.0675 | -104.9037
16.0713 | -104.9027
16.075 | -104.9007
16.0788 | -104.9003
16.0825 | -104.9032
16.0863 | -104.9071
16.09 | -104.9097
16.0938 | -104.9091
16.0975 | -104.9086
16.1013 | -104.9061
16.105 | -104.9011
16.1088 | -104.8964
16.1125 | -104.8951
16.1163 | -104.8939
16.12 | -104.8907
16.1238 | -104.8845
16.1275 | -104.8791
16.1313 | -104.8704
16.135 | -104.8638
16.1388 | -104.8568
16.1425 | -104.8479
16.1463 | -104.8374
16.15 | -104.8279
16.1538 | -104.8206
16.1575 | -104.8127
16.1613 | -104.8046
16.165 | -104.7928
16.1688 | -104.7828
16.1725 | -104.7768
16.1763 | -104.7646
16.18 | -104.7553
16.1838 | -104.7499
16.1875 | -104.746
16.1913 | -104.7419
16.195 | -104.7347
16.1988 | -104.7293
16.2025 | -104.7291
16.2063 | -104.7301
16.21 | -104.7328
16.2138 | -104.7332
16.2175 | -104.7365
16.2213 | -104.7382
16.225 | -104.7389
16.2288 | -104.7417
16.2325 | -104.7446
16.2363 | -104.7492
16.24 | -104.7561
16.2438 | -104.7635
16.2475 | -104.7722
16.2513 | -104.779
16.255 | -104.7846
16.2588 | -104.7897
16.2625 | -104.7909
16.2663 | -104.7952
16.27 | -104.8017
16.2738 | -104.8047
16.2775 | -104.8068
16.2813 | -104.8115
16.285 | -104.8181
16.2888 | -104.8214
16.2925 | -104.8256
16.2963 | -104.8284
16.3 | -104.8318
16.3038 | -104.838
16.3075 | -104.8444
16.3113 | -104.8497
16.315 | -104.8515
16.3188 | -104.8527
16.3225 | -104.8535
16.3263 | -104.8586
16.3301 | -104.8599
16.3338 | -104.859
16.3376 | -104.8602
16.3413 | -104.8604
16.3451 | -104.8583
16.3488 | -104.8556
16.3526 | -104.8525
16.3563 | -104.8516
16.3601 | -104.8485
16.3638 | -104.8469
16.3676 | -104.8431
16.3713 | -104.8393
16.3751 | -104.8373
16.3788 | -104.8316
16.3826 | -104.8234
16.3863 | -104.8154
16.3901 | -104.8089
16.3938 | -104.8001
16.3976 | -104.7934
16.4013 | -104.7837
16.4051 | -104.7745
16.4088 | -104.7679
16.4126 | -104.7599
16.4163 | -104.75
16.4201 | -104.7405
16.4238 | -104.7312
16.4276 | -104.7233
16.4313 | -104.7161
16.4351 | -104.7088
16.4388 | -104.7027
16.4426 | -104.6991
16.4463 | -104.6996
16.4501 | -104.6975
16.4538 | -104.6984
16.4576 | -104.6998
16.4613 | -104.7007
16.4651 | -104.7033
16.4688 | -104.7056
16.4726 | -104.7077
16.4763 | -104.7109
16.4801 | -104.7133
16.4838 | -104.7154
16.4876 | -104.7211
16.4913 | -104.729
16.4951 | -104.7373
16.4988 | -104.7458
16.5026 | -104.7529
16.5063 | -104.7605
16.5101 | -104.7709
16.5138 | -104.7794
16.5176 | -104.7859
16.5213 | -104.7929
16.5251 | -104.7992
16.5288 | -104.8051
16.5326 | -104.8132
16.5363 | -104.8181
16.5401 | -104.8223
16.5438 | -104.8289
16.5476 | -104.8364
16.5513 | -104.8426
16.5551 | -104.8499
16.5588 | -104.8573
16.5626 | -104.861
16.5663 | -104.8656
16.5701 | -104.8709
16.5738 | -104.8723
16.5776 | -104.8735
16.5813 | -104.8757
16.5851 | -104.8764
16.5888 | -104.877
16.5926 | -104.8753
16.5963 | -104.8734
16.6001 | -104.8702
16.6038 | -104.8658
16.6076 | -104.8594
16.6113 | -104.8524
16.6151 | -104.8483
16.6188 | -104.843
16.6226 | -104.8354
16.6263 | -104.8239
16.6301 | -104.8158
16.6338 | -104.809
16.6376 | -104.8018
16.6413 | -104.7911
16.6451 | -104.7833
16.6489 | -104.775
16.6526 | -104.765
16.6564 | -104.7586
16.6601 | -104.7509
16.6639 | -104.7411
16.6676 | -104.7322
16.6714 | -104.7232
16.6751 | -104.7143
16.6789 | -104.7044
16.6826 | -104.693
16.6864 | -104.6861
16.6901 | -104.6771
16.6939 | -104.6677
16.6976 | -104.6594
16.7014 | -104.6515
16.7051 | -104.6435
16.7089 | -104.6342
16.7126 | -104.6234
16.7164 | -104.6126
16.7201 | -104.6023
16.7239 | -104.5923
16.7276 | -104.5818
16.7314 | -104.5714
16.7351 | -104.5615
16.7389 | -104.5528
16.7426 | -104.5462
16.7464 | -104.5385
16.7501 | -104.532
16.7539 | -104.5256
16.7576 | -104.5181
16.7614 | -104.5132
16.7651 | -104.5078
16.7689 | -104.5028
16.7726 | -104.5037
16.7764 | -104.5045
16.7801 | -104.5027
16.7839 | -104.499
16.7876 | -104.4948
16.7914 | -104.4942
16.7951 | -104.4938
16.7989 | -104.4908
16.8026 | -104.4895
16.8064 | -104.4887
16.8101 | -104.4897
16.8139 | -104.4912
16.8176 | -104.4911
16.8214 | -104.49
16.8251 | -104.4898
16.8289 | -104.4901
16.8326 | -104.4892
16.8364 | -104.4889
16.8401 | -104.4895
16.8439 | -104.4894
16.8476 | -104.4881
16.8514 | -104.486
16.8551 | -104.4826
16.8589 | -104.4775
16.8626 | -104.4704
16.8664 | -104.4647
16.8701 | -104.4596
16.8739 | -104.453
16.8776 | -104.4469
16.8814 | -104.4414
16.8851 | -104.4339
16.8889 | -104.4266
16.8926 | -104.4195
16.8964 | -104.4082
16.9001 | -104.3974
16.9039 | -104.3867
16.9076 | -104.3748
16.9114 | -104.3614
16.9151 | -104.3481
16.9189 | -104.3358
16.9226 | -104.325
16.9264 | -104.313
16.9301 | -104.3028
16.9339 | -104.2933
16.9376 | -104.2834
16.9414 | -104.2728
16.9451 | -104.2647
16.9489 | -104.2566
16.9526 | -104.247
16.9564 | -104.2389
16.9601 | -104.2322
16.9639 | -104.2283
16.9677 | -104.224
16.9714 | -104.2204
16.9752 | -104.2185
16.9789 | -104.2159
16.9827 | -104.2121
16.9864 | -104.2099
16.9902 | -104.2101
16.9939 | -104.2123
16.9977 | -104.2131
17.0014 | -104.2129
17.0052 | -104.2147
17.0089 | -104.2166
17.0127 | -104.218
17.0164 | -104.2204
17.0202 | -104.2192
17.0239 | -104.2187
17.0277 | -104.2206
17.0314 | -104.2215
17.0352 | -104.2222
17.0389 | -104.2209
17.0427 | -104.2224
17.0464 | -104.2219
17.0502 | -104.2189
17.0539 | -104.2155
17.0577 | -104.2117
17.0614 | -104.2053
17.0652 | -104.1991
17.0689 | -104.1938
17.0727 | -104.189
17.0764 | -104.184
17.0802 | -104.1786
17.0839 | -104.1728
17.0877 | -104.1637
17.0914 | -104.1565
17.0952 | -104.1506
17.0989 | -104.1445
17.1027 | -104.1416
17.1064 | -104.1407
17.1102 | -104.1367
17.1139 | -104.1299
17.1177 | -104.1238
17.1214 | -104.1177
17.1252 | -104.1107
17.1289 | -104.1039
17.1327 | -104.1015
17.1364 | -104.1015
17.1402 | -104.1035
17.1439 | -104.1039
17.1477 | -104.1057
17.1514 | -104.1079
17.1552 | -104.1102
17.1589 | -104.1116
17.1627 | -104.1129
17.1664 | -104.1139
17.1702 | -104.1147
17.1739 | -104.1154
17.1777 | -104.1179
17.1814 | -104.1188
17.1852 | -104.1189
17.1889 | -104.1208
17.1927 | -104.1229
17.1964 | -104.1244
17.2002 | -104.1247
17.2039 | -104.1235
17.2077 | -104.1197
17.2114 | -104.1164
17.2152 | -104.1143
17.2189 | -104.1107
17.2227 | -104.1055
17.2264 | -104.1009
17.2302 | -104.0949
17.2339 | -104.0857
17.2377 | -104.0789
17.2414 | -104.0752
17.2452 | -104.0679
17.2489 | -104.0588
17.2527 | -104.0503
17.2564 | -104.0393
17.2602 | -104.0261
17.2639 | -104.0172
17.2677 | -104.0099
17.2714 | -103.9991
17.2752 | -103.988
17.2789 | -103.9753
17.2827 | -103.9622
17.2865 | -103.9516
17.2902 | -103.9424
17.294 | -103.9339
17.2977 | -103.9233
17.3015 | -103.9151
17.3052 | -103.9106
17.309 | -103.9064
17.3127 | -103.9047
17.3165 | -103.8988
17.3202 | -103.8926
17.324 | -103.8849
17.3277 | -103.8778
17.3315 | -103.8734
17.3352 | -103.8727
17.339 | -103.8703
17.3427 | -103.8671
17.3465 | -103.8652
17.3502 | -103.8616
17.354 | -103.858
17.3577 | -103.8575
17.3615 | -103.856
17.3652 | -103.8559
17.369 | -103.857
17.3727 | -103.8603
17.3765 | -103.8631
17.3802 | -103.8623
17.384 | -103.8624
17.3877 | -103.8634
17.3915 | -103.8656
17.3952 | -103.8668
17.399 | -103.8676
17.4027 | -103.8695
17.4065 | -103.8712
17.4102 | -103.8716
17.414 | -103.8718
17.4177 | -103.8721
17.4215 | -103.8729
17.4252 | -103.8727
17.429 | -103.8722
17.4327 | -103.8717
17.4365 | -103.8693
17.4402 | -103.8686
17.444 | -103.8701
17.4477 | -103.8696
17.4515 | -103.8682
17.4552 | -103.8673
17.459 | -103.8659
17.4627 | -103.8631
17.4665 | -103.8581
17.4702 | -103.8565
17.474 | -103.854
17.4777 | -103.8506
17.4815 | -103.8471
17.4852 | -103.8429
17.489 | -103.8377
17.4927 | -103.8355
17.4965 | -103.8334
17.5002 | -103.8292
17.504 | -103.8243
17.5077 | -103.8189
17.5115 | -103.8145
17.5152 | -103.8086
17.519 | -103.8027
17.5227 | -103.7965
17.5265 | -103.7896
17.5302 | -103.7853
17.534 | -103.7772
17.5377 | -103.7664
17.5415 | -103.757
17.5452 | -103.7487
17.549 | -103.7398
17.5527 | -103.7292
17.5565 | -103.7185
17.5602 | -103.708
17.564 | -103.7001
17.5677 | -103.692
17.5715 | -103.6834
17.5752 | -103.6765
17.579 | -103.6703
17.5827 | -103.6619
17.5865 | -103.6541
17.5902 | -103.6479
17.594 | -103.6401
17.5978 | -103.6329
17.6015 | -103.6282
17.6053 | -103.6237
17.609 | -103.6206
17.6128 | -103.6167
17.6165 | -103.6111
17.6203 | -103.6073
17.624 | -103.6031
17.6278 | -103.5994
17.6315 | -103.5965
17.6353 | -103.5937
17.639 | -103.5892
17.6428 | -103.5864
17.6465 | -103.5839
17.6503 | -103.5807
17.654 | -103.5792
17.6578 | -103.5789
17.6615 | -103.5774
17.6653 | -103.5757
17.669 | -103.5733
17.6728 | -103.5692
17.6765 | -103.5654
17.6803 | -103.5618
17.684 | -103.5587
17.6878 | -103.5512
17.6915 | -103.5428
17.6953 | -103.5354
17.699 | -103.5313
17.7028 | -103.527
17.7065 | -103.5213
17.7103 | -103.5171
17.714 | -103.5121
17.7178 | -103.5093
17.7215 | -103.5054
17.7253 | -103.5036
17.729 | -103.5031
17.7328 | -103.505
17.7365 | -103.5021
17.7403 | -103.4993
17.744 | -103.5004
17.7478 | -103.4992
17.7515 | -103.5021
17.7553 | -103.5092
17.759 | -103.5161
17.7628 | -103.5209
17.7665 | -103.5268
17.7703 | -103.5341
17.774 | -103.5418
17.7778 | -103.5498
17.7815 | -103.5561
17.7853 | -103.5627
17.789 | -103.5733
17.7928 | -103.5844
17.7965 | -103.594
17.8003 | -103.6005
17.804 | -103.6079
17.8078 | -103.6146
17.8115 | -103.619
17.8153 | -103.6221
17.819 | -103.6267
17.8228 | -103.6316
17.8265 | -103.637
17.8303 | -103.6397
17.834 | -103.6415
17.8378 | -103.6433
17.8415 | -103.6474
17.8453 | -103.6521
17.849 | -103.6553
17.8528 | -103.6515
17.8565 | -103.6521
17.8603 | -103.6516
17.864 | -103.6487
17.8678 | -103.6489
17.8715 | -103.6473
17.8753 | -103.6447
17.879 | -103.6437
17.8828 | -103.6418
17.8865 | -103.6386
17.8903 | -103.6346
17.894 | -103.631
17.8978 | -103.6278
17.9015 | -103.6219
17.9053 | -103.6177
17.909 | -103.6118
17.9128 | -103.6065
17.9165 | -103.6038
17.9203 | -103.5998
17.9241 | -103.5965
17.9278 | -103.5935
17.9316 | -103.5908
17.9353 | -103.5871
17.9391 | -103.5834
17.9428 | -103.5843
17.9466 | -103.5818
17.9503 | -103.5793
17.9541 | -103.5766
17.9578 | -103.5734
17.9616 | -103.5695
17.9653 | -103.5633
17.9691 | -103.5604
17.9728 | -103.5578
17.9766 | -103.5544
17.9803 | -103.5465
17.9841 | -103.5378
17.9878 | -103.5271
17.9916 | -103.5198
17.9953 | -103.5135
17.9991 | -103.5041
18.0028 | -103.4955
18.0066 | -103.4836
18.0103 | -103.4704
18.0141 | -103.4593
18.0178 | -103.447
18.0216 | -103.4338
18.0253 | -103.4198
18.0291 | -103.4053
18.0328 | -103.3932
18.0366 | -103.3792
18.0403 | -103.3622
18.0441 | -103.3446
18.0478 | -103.3273
18.0516 | -103.3099
18.0553 | -103.2967
18.0591 | -103.2841
18.0628 | -103.2727
18.0666 | -103.2626
18.0703 | -103.2526
18.0741 | -103.2395
18.0778 | -103.2251
18.0816 | -103.2111
18.0853 | -103.1988
18.0891 | -103.1894
18.0928 | -103.1785
18.0966 | -103.168
18.1003 | -103.1574
18.1041 | -103.1487
18.1078 | -103.1425
18.1116 | -103.1376
18.1153 | -103.1333
18.1191 | -103.1307
18.1228 | -103.1279
18.1266 | -103.1246
18.1303 | -103.1221
18.1341 | -103.1185
18.1378 | -103.1155
18.1416 | -103.1139
18.1453 | -103.1146
18.1491 | -103.1146
18.1528 | -103.1128
18.1566 | -103.1111
18.1603 | -103.1111
18.1641 | -103.1097
18.1678 | -103.107
18.1716 | -103.1016
18.1753 | -103.0945
18.1791 | -103.0912
18.1828 | -103.0862
18.1866 | -103.079
18.1903 | -103.0705
18.1941 | -103.0632
18.1978 | -103.0545
18.2016 | -103.0432
18.2053 | -103.0308
18.2091 | -103.0161
18.2128 | -102.9997
18.2166 | -102.9806
18.2203 | -102.9608
18.2241 | -102.9441
18.2278 | -102.9232
18.2316 | -102.9028
18.2353 | -102.8807
18.2391 | -102.8577
18.2429 | -102.8356
18.2466 | -102.812
18.2504 | -102.7862
18.2541 | -102.7594
18.2579 | -102.7325
18.2616 | -102.7082
18.2654 | -102.683
18.2691 | -102.6584
18.2729 | -102.6336
18.2766 | -102.6096
18.2804 | -102.5864
18.2841 | -102.5655
18.2879 | -102.547
18.2916 | -102.5267
18.2954 | -102.5083
18.2991 | -102.4941
18.3029 | -102.4803
18.3066 | -102.4667
18.3104 | -102.4573
18.3141 | -102.4475
18.3179 | -102.4379
18.3216 | -102.4293
18.3254 | -102.422
18.3291 | -102.4166
18.3329 | -102.4158
18.3366 | -102.4172
18.3404 | -102.4191
18.3441 | -102.4218
18.3479 | -102.4236
18.3516 | -102.4273
18.3554 | -102.4298
18.3591 | -102.4335
18.3629 | -102.4354
18.3666 | -102.4389
18.3704 | -102.4428
18.3741 | -102.4442
18.3779 | -102.4447
18.3816 | -102.4494
18.3854 | -102.4515
18.3891 | -102.4526
18.3929 | -102.4547
18.3966 | -102.449
18.4004 | -102.4478
18.4041 | -102.4448
18.4079 | -102.4387
18.4116 | -102.4325
18.4154 | -102.4282
18.4191 | -102.421
18.4229 | -102.4116
18.4266 | -102.4031
18.4304 | -102.395
18.4341 | -102.3866
18.4379 | -102.3782
18.4416 | -102.3647
18.4454 | -102.3518
18.4491 | -102.3406
18.4529 | -102.3272
18.4566 | -102.3128
18.4604 | -102.2986
18.4641 | -102.2843
18.4679 | -102.2693
18.4716 | -102.2531
18.4754 | -102.2372
18.4791 | -102.2223
18.4829 | -102.2093
18.4866 | -102.1925
18.4904 | -102.1777
18.4941 | -102.1622
18.4979 | -102.1483
18.5016 | -102.1375
18.5054 | -102.1268
18.5091 | -102.1162
18.5129 | -102.1038
18.5166 | -102.0917
18.5204 | -102.0802
18.5241 | -102.0696
18.5279 | -102.062
18.5316 | -102.0541
18.5354 | -102.0453
18.5391 | -102.0367
18.5429 | -102.0275
18.5466 | -102.0196
18.5504 | -102.0131
18.5541 | -102.0061
18.5579 | -101.9991
18.5617 | -101.991
18.5654 | -101.9821
18.5692 | -101.973
18.5729 | -101.9609
18.5767 | -101.9474
18.5804 | -101.9355
18.5842 | -101.9233
18.5879 | -101.9116
18.5917 | -101.8994
18.5954 | -101.8838
18.5992 | -101.868
18.6029 | -101.8545
18.6067 | -101.8401
18.6104 | -101.8237
18.6142 | -101.8077
18.6179 | -101.7929
18.6217 | -101.7774
18.6254 | -101.7622
18.6292 | -101.7466
18.6329 | -101.7289
18.6367 | -101.7137
18.6404 | -101.6956
18.6442 | -101.6727
18.6479 | -101.6537
18.6517 | -101.6363
18.6554 | -101.618
18.6592 | -101.6008
18.6629 | -101.5833
18.6667 | -101.5667
18.6704 | -101.5498
18.6742 | -101.535
18.6779 | -101.5222
18.6817 | -101.5092
18.6854 | -101.4961
18.6892 | -101.4853
18.6929 | -101.4745
18.6967 | -101.4622
18.7004 | -101.4518
18.7042 | -101.4434
18.7079 | -101.439
18.7117 | -101.434
18.7154 | -101.4267
18.7192 | -101.421
18.7229 | -101.4197
18.7267 | -101.4191
18.7304 | -101.4198
18.7342 | -101.4213
18.7379 | -101.4239
18.7417 | -101.4258
18.7454 | -101.426
18.7492 | -101.4246
18.7529 | -101.4275
18.7567 | -101.4297
18.7604 | -101.4316
18.7642 | -101.4336
18.7679 | -101.4339
18.7717 | -101.434
18.7754 | -101.436
18.7792 | -101.4375
18.7829 | -101.4385
18.7867 | -101.4395
18.7904 | -101.4363
18.7942 | -101.4321
18.7979 | -101.4283
18.8017 | -101.4272
18.8054 | -101.426
18.8092 | -101.4245
18.8129 | -101.4205
18.8167 | -101.4169
18.8204 | -101.4139
18.8242 | -101.41
18.8279 | -101.4052
18.8317 | -101.3979
18.8354 | -101.392
18.8392 | -101.3886
18.8429 | -101.3844
18.8467 | -101.3776
18.8504 | -101.3691
18.8542 | -101.3622
18.8579 | -101.3538
18.8617 | -101.3446
18.8654 | -101.3353
18.8692 | -101.3223
18.8729 | -101.3121
18.8767 | -101.3039
18.8804 | -101.2943
18.8842 | -101.2838
18.888 | -101.2738
18.8917 | -101.2642
18.8955 | -101.253
18.8992 | -101.2422
18.903 | -101.2333
18.9067 | -101.2258
18.9105 | -101.2201
18.9142 | -101.2117
18.918 | -101.2033
18.9217 | -101.1958
18.9255 | -101.1871
18.9292 | -101.1764
18.933 | -101.1667
18.9367 | -101.1651
18.9405 | -101.1593
18.9442 | -101.1548
18.948 | -101.1497
18.9517 | -101.1467
18.9555 | -101.1457
18.9592 | -101.1441
18.963 | -101.1425
18.9667 | -101.1454
18.9705 | -101.1473
18.9742 | -101.1475
18.978 | -101.1485
18.9817 | -101.1532
18.9855 | -101.1595
18.9892 | -101.1648
18.993 | -101.1681
18.9967 | -101.1731
19.0005 | -101.179
19.0042 | -101.186
19.008 | -101.1921
19.0117 | -101.1981
19.0155 | -101.2065
19.0192 | -101.2163
19.023 | -101.223
19.0267 | -101.2309
19.0305 | -101.2372
19.0342 | -101.243
19.038 | -101.2503
19.0417 | -101.2564
19.0455 | -101.2609
19.0492 | -101.267
19.053 | -101.271
19.0567 | -101.2707
19.0605 | -101.2712
19.0642 | -101.2721
19.068 | -101.2728
19.0717 | -101.2726
19.0755 | -101.2706
19.0792 | -101.2679
19.083 | -101.2653
19.0867 | -101.2619
19.0905 | -101.257
19.0942 | -101.2523
19.098 | -101.2488
19.1017 | -101.246
19.1055 | -101.2411
19.1092 | -101.2341
19.113 | -101.2247
19.1167 | -101.2179
19.1205 | -101.2129
19.1242 | -101.2048
19.128 | -101.1977
19.1317 | -101.1922
19.1355 | -101.1862
19.1392 | -101.1792
19.143 | -101.1703
19.1467 | -101.162
19.1505 | -101.1537
19.1542 | -101.1451
19.158 | -101.1383
19.1617 | -101.1325
19.1655 | -101.1258
19.1692 | -101.1184
19.173 | -101.1111
19.1767 | -101.1046
19.1805 | -101.097
19.1842 | -101.0886
19.188 | -101.0807
19.1917 | -101.0712
19.1955 | -101.0635
19.1992 | -101.059
19.203 | -101.0515
19.2068 | -101.0414
19.2105 | -101.0322
19.2143 | -101.0249
19.218 | -101.0149
19.2218 | -101.0033
19.2255 | -100.9919
19.2293 | -100.9817
19.233 | -100.9727
19.2368 | -100.9634
19.2405 | -100.9529
19.2443 | -100.9446
19.248 | -100.9365
19.2518 | -100.9278
19.2555 | -100.9189
19.2593 | -100.9121
19.263 | -100.9067
19.2668 | -100.8979
19.2705 | -100.8902
19.2743 | -100.8828
19.278 | -100.8747
19.2818 | -100.8675
19.2855 | -100.8623
19.2893 | -100.8564
19.293 | -100.8504
19.2968 | -100.8453
19.3005 | -100.8401
19.3043 | -100.836
19.308 | -100.8316
19.3118 | -100.8263
19.3155 | -100.82
19.3193 | -100.8143
19.323 | -100.8105
19.3268 | -100.8088
19.3305 | -100.8096
19.3343 | -100.8083
19.338 | -100.805
19.3418 | -100.8007
19.3455 | -100.7921
19.3493 | -100.7844
19.353 | -100.7782
19.3568 | -100.7721
19.3605 | -100.7656
19.3643 | -100.761
19.368 | -100.7567
19.3718 | -100.7495
19.3755 | -100.7443
19.3793 | -100.7375
19.383 | -100.73
19.3868 | -100.7226
19.3905 | -100.7135
19.3943 | -100.7044
19.398 | -100.6959
19.4018 | -100.6861
19.4055 | -100.6766
19.4093 | -100.6665
19.413 | -100.6582
19.4168 | -100.6501
19.4205 | -100.6401
19.4243 | -100.6306
19.428 | -100.6208
19.4318 | -100.6078
19.4355 | -100.595
19.4393 | -100.586
19.443 | -100.5797
19.4468 | -100.5699
19.4505 | -100.56
19.4543 | -100.5536
19.458 | -100.5503
19.4618 | -100.546
19.4655 | -100.5422
19.4693 | -100.5384
19.473 | -100.5343
19.4768 | -100.5362
19.4805 | -100.5358
19.4843 | -100.5345
19.488 | -100.5311
19.4918 | -100.5294
19.4955 | -100.5279
19.4993 | -100.5277
19.503 | -100.5296
19.5068 | -100.5312
19.5105 | -100.5345
19.5143 | -100.5389
19.518 | -100.5431
19.5218 | -100.5464
19.5256 | -100.5501
19.5293 | -100.554
19.5331 | -100.5586
19.5368 | -100.5636
19.5406 | -100.566
19.5443 | -100.5692
19.5481 | -100.5729
19.5518 | -100.5745
19.5556 | -100.5768
19.5593 | -100.5803
19.5631 | -100.583
19.5668 | -100.5858
19.5706 | -100.5863
19.5743 | -100.5861
19.5781 | -100.5856
19.5818 | -100.5837
19.5856 | -100.5822
19.5893 | -100.5816
19.5931 | -100.58
19.5968 | -100.5778
19.6006 | -100.5719
19.6043 | -100.5633
19.6081 | -100.5565
19.6118 | -100.5519
19.6156 | -100.5453
19.6193 | -100.5366
19.6231 | -100.5285
19.6268 | -100.519
19.6306 | -100.5085
19.6343 | -100.4989
19.6381 | -100.4891
19.6418 | -100.4798
19.6456 | -100.47
19.6493 | -100.4599
19.6531 | -100.4476
19.6568 | -100.4363
19.6606 | -100.4251
19.6643 | -100.4167
19.6681 | -100.405
19.6718 | -100.3907
19.6756 | -100.3759
19.6793 | -100.3627
19.6831 | -100.3493
19.6868 | -100.3355
19.6906 | -100.322
19.6943 | -100.3086
19.6981 | -100.2967
19.7018 | -100.2818
19.7056 | -100.2676
19.7093 | -100.2574
19.7131 | -100.2429
19.7168 | -100.2247
19.7206 | -100.2065
19.7243 | -100.1907
19.7281 | -100.1785
19.7318 | -100.1627
19.7356 | -100.1468
19.7393 | -100.1326
19.7431 | -100.12
19.7468 | -100.1085
19.7506 | -100.0998
19.7543 | -100.0911
19.7581 | -100.0806
19.7618 | -100.0699
19.7656 | -100.0603
19.7693 | -100.0516
19.7731 | -100.0429
19.7768 | -100.0373
19.7806 | -100.0321
19.7843 | -100.0239
19.7881 | -100.0181
19.7918 | -100.0114
19.7956 | -100.0041
19.7993 | -99.9984
19.8031 | -99.9936
19.8068 | -99.9885
19.8106 | -99.9843
19.8143 | -99.9803
19.8181 | -99.9764
19.8218 | -99.9756
19.8256 | -99.9754
19.8293 | -99.9739
19.8331 | -99.9738
19.8368 | -99.9736
19.8406 | -99.9723
19.8444 | -99.9707
19.8481 | -99.9688
19.8519 | -99.9691
19.8556 | -99.9726
19.8594 | -99.9775
19.8631 | -99.9851
19.8669 | -99.9911
19.8706 | -99.9974
19.8744 | -100.0074
19.8781 | -100.0163
19.8819 | -100.027
19.8856 | -100.0387
19.8894 | -100.0492
19.8931 | -100.0577
19.8969 | -100.069
19.9006 | -100.0837
19.9044 | -100.0988
19.9081 | -100.1132
19.9119 | -100.1268
19.9156 | -100.1384
19.9194 | -100.1489
19.9231 | -100.159
19.9269 | -100.1699
19.9306 | -100.1813
19.9344 | -100.1925
19.9381 | -100.2042
19.9419 | -100.2149
19.9456 | -100.224
19.9494 | -100.2326
19.9531 | -100.2423
19.9569 | -100.2515
19.9606 | -100.2575
19.9644 | -100.2592
19.9681 | -100.2611
19.9719 | -100.2636
19.9756 | -100.2658
19.9794 | -100.2673
19.9831 | -100.2704
19.9869 | -100.2689
19.9906 | -100.2632
19.9944 | -100.262
19.9981 | -100.2623
20.0019 | -100.2608
20.0056 | -100.2567
20.0094 | -100.2477
20.0131 | -100.2382
20.0169 | -100.2172
20.0206 | -100.208
20.0244 | -100.1966
20.0281 | -100.182
20.0319 | -100.1664
20.0356 | -100.1466
20.0394 | -100.1305
20.0431 | -100.1149
20.0469 | -100.0997
20.0506 | -100.0856
20.0544 | -100.0711
20.0581 | -100.0548
20.0619 | -100.0357
20.0656 | -100.0156
20.0694 | -99.9976
20.0731 | -99.9786
20.0769 | -99.9603
20.0806 | -99.943
20.0844 | -99.9269
20.0881 | -99.9122
20.0919 | -99.8998
20.0956 | -99.8857
20.0994 | -99.8715
20.1031 | -99.8595
20.1069 | -99.8478
20.1106 | -99.8364
20.1144 | -99.827
20.1181 | -99.8195
20.1219 | -99.8121
20.1256 | -99.8062
20.1294 | -99.8018
20.1331 | -99.7986
20.1369 | -99.7955
20.1406 | -99.7906
20.1444 | -99.7859
20.1481 | -99.7827
20.1519 | -99.7833
20.1556 | -99.7845
20.1594 | -99.7851
20.1632 | -99.7864
20.1669 | -99.7894
20.1707 | -99.792
20.1744 | -99.7957
20.1782 | -99.8004
20.1819 | -99.8049
20.1857 | -99.808
20.1894 | -99.8122
20.1932 | -99.8179
20.1969 | -99.8227
20.2007 | -99.8268
20.2044 | -99.8315
20.2082 | -99.8358
20.2119 | -99.8381
20.2157 | -99.8417
20.2194 | -99.8454
20.2232 | -99.8461
20.2269 | -99.8444
20.2307 | -99.8431
20.2344 | -99.8447
20.2382 | -99.8458
20.2419 | -99.8439
20.2457 | -99.8415
20.2494 | -99.8415
20.2532 | -99.8404
20.2569 | -99.8348
20.2607 | -99.8296
20.2644 | -99.8251
20.2682 | -99.8194
20.2719 | -99.8151
20.2757 | -99.8108
20.2794 | -99.8047
20.2832 | -99.7981
20.2869 | -99.7911
20.2907 | -99.7848
20.2944 | -99.7804
20.2982 | -99.7762
20.3019 | -99.7732
20.3057 | -99.7676
20.3094 | -99.7597
20.3132 | -99.7545
20.3169 | -99.7521
20.3207 | -99.7507
20.3244 | -99.7496
20.3282 | -99.7465
20.3319 | -99.7429
20.3357 | -99.7406
20.3394 | -99.7383
20.3432 | -99.7367
20.3469 | -99.7355
20.3507 | -99.7346
20.3544 | -99.7331
20.3582 | -99.7293
20.3619 | -99.7246
20.3657 | -99.7226
20.3694 | -99.7229
20.3732 | -99.7219
20.3769 | -99.7189
20.3807 | -99.7147
20.3844 | -99.7134
20.3882 | -99.7116
20.3919 | -99.7085
20.3957 | -99.7054
20.3994 | -99.7019
20.4032 | -99.6967
20.4069 | -99.691
20.4107 | -99.6859
20.4144 | -99.6789
20.4182 | -99.6687
20.4219 | -99.6604
20.4257 | -99.6534
20.4294 | -99.6458
20.4332 | -99.6376
20.4369 | -99.6269
20.4407 | -99.616
20.4444 | -99.6058
20.4482 | -99.5934
20.4519 | -99.5825
20.4557 | -99.5703
20.4594 | -99.5598
20.4632 | -99.5488
20.4669 | -99.5394
20.4707 | -99.5282
20.4744 | -99.5193
20.4782 | -99.5108
20.482 | -99.501
20.4857 | -99.4921
20.4895 | -99.4832
20.4932 | -99.4741
20.497 | -99.466
20.5007 | -99.4605
20.5045 | -99.4544
20.5082 | -99.4479
20.512 | -99.4452
20.5157 | -99.445
20.5195 | -99.4445
20.5232 | -99.4421
20.527 | -99.4397
20.5307 | -99.4408
20.5345 | -99.4428
20.5382 | -99.4431
20.542 | -99.4434
20.5457 | -99.4463
20.5495 | -99.4485
20.5532 | -99.4511
20.557 | -99.451
20.5607 | -99.4501
20.5645 | -99.452
20.5682 | -99.4557
20.572 | -99.4575
20.5757 | -99.4558
20.5795 | -99.4569
20.5832 | -99.4592
20.587 | -99.46
20.5907 | -99.4592
20.5945 | -99.457
20.5982 | -99.4556
20.602 | -99.4529
20.6057 | -99.447
20.6095 | -99.4405
20.6132 | -99.4328
20.617 | -99.426
20.6207 | -99.4179
20.6245 | -99.4085
20.6282 | -99.3988
20.632 | -99.3898
20.6357 | -99.3777
20.6395 | -99.3644
20.6432 | -99.3526
20.647 | -99.3439
20.6507 | -99.3351
20.6545 | -99.3267
20.6582 | -99.3136
20.662 | -99.2964
20.6657 | -99.2789
20.6695 | -99.262
20.6732 | -99.2465
20.677 | -99.2285
20.6807 | -99.2111
20.6845 | -99.1936
20.6882 | -99.1772
20.692 | -99.1624
20.6957 | -99.1477
20.6995 | -99.1308
20.7032 | -99.1159
20.707 | -99.1018
20.7107 | -99.0893
20.7145 | -99.0758
20.7182 | -99.0614
20.722 | -99.0481
20.7257 | -99.037
20.7295 | -99.0272
20.7332 | -99.017
20.737 | -99.0046
20.7407 | -98.9953
20.7445 | -98.9872
20.7482 | -98.9794
20.752 | -98.9734
20.7557 | -98.9687
20.7595 | -98.9649
20.7632 | -98.9635
20.767 | -98.9611
20.7707 | -98.9586
20.7745 | -98.9561
20.7782 | -98.9585
20.782 | -98.9624
20.7857 | -98.9656
20.7895 | -98.9692
20.7932 | -98.9725
20.797 | -98.9767
20.8008 | -98.9823
20.8045 | -98.9871
20.8083 | -98.9908
20.812 | -98.9931
20.8158 | -98.9974
20.8195 | -99.0014
20.8233 | -99.0036
20.827 | -99.0067
20.8308 | -99.0093
20.8345 | -99.0126
20.8383 | -99.0168
20.842 | -99.0192
20.8458 | -99.0202
20.8495 | -99.0264
20.8533 | -99.0288
20.857 | -99.0302
20.8608 | -99.0309
20.8645 | -99.0296
20.8683 | -99.0286
20.872 | -99.0294
20.8758 | -99.0265
20.8795 | -99.0231
20.8833 | -99.0202
20.887 | -99.0191
20.8908 | -99.018
20.8945 | -99.0155
20.8983 | -99.0142
20.902 | -99.0141
20.9058 | -99.0144
20.9095 | -99.0136
20.9133 | -99.0114
20.917 | -99.0097
20.9208 | -99.0065
20.9245 | -99.0012
20.9283 | -98.9961
20.932 | -98.9949
20.9358 | -98.9942
20.9395 | -98.9914
20.9433 | -98.9868
20.947 | -98.9816
20.9508 | -98.9754
20.9545 | -98.972
20.9583 | -98.9715
20.962 | -98.9686
20.9658 | -98.967
20.9695 | -98.9638
20.9733 | -98.9622
20.977 | -98.9592
20.9808 | -98.956
20.9845 | -98.952
20.9883 | -98.9465
20.992 | -98.9401
20.9958 | -98.9358
20.9995 | -98.9314
21.0033 | -98.9265
21.007 | -98.9229
21.0108 | -98.9194
21.0145 | -98.915
21.0183 | -98.911
21.022 | -98.9043
21.0258 | -98.8996
21.0295 | -98.8935
21.0333 | -98.8898
21.037 | -98.8865
21.0408 | -98.8808
21.0445 | -98.8766
21.0483 | -98.8709
21.052 | -98.8649
21.0558 | -98.8586
21.0595 | -98.8527
21.0633 | -98.8467
21.067 | -98.8425
21.0708 | -98.8407
21.0745 | -98.8354
21.0783 | -98.8309
21.082 | -98.8278
21.0858 | -98.8222
21.0895 | -98.8159
21.0933 | -98.8108
21.097 | -98.8051
21.1008 | -98.8059
21.1045 | -98.803
21.1083 | -98.8004
21.112 | -98.8004
21.1158 | -98.8005
21.1196 | -98.8014
21.1233 | -98.8036
21.1271 | -98.8059
21.1308 | -98.8082
21.1346 | -98.8105
21.1383 | -98.8125
21.1421 | -98.8158
21.1458 | -98.8192
21.1496 | -98.8219
21.1533 | -98.8262
21.1571 | -98.8291
21.1608 | -98.8298
21.1646 | -98.8311
21.1683 | -98.8349
21.1721 | -98.8396
21.1758 | -98.8427
21.1796 | -98.8439
21.1833 | -98.8461
21.1871 | -98.8467
21.1908 | -98.8484
21.1946 | -98.8508
21.1983 | -98.8513
21.2021 | -98.8493
21.2058 | -98.8445
21.2096 | -98.8406
21.2133 | -98.8382
21.2171 | -98.8342
21.2208 | -98.8288
21.2246 | -98.8247
21.2283 | -98.8185
21.2321 | -98.8096
21.2358 | -98.8007
21.2396 | -98.7889
21.2433 | -98.7753
21.2471 | -98.7607
21.2508 | -98.745
21.2546 | -98.7283
21.2583 | -98.7102
21.2621 | -98.6919
21.2658 | -98.6719
21.2696 | -98.6527
21.2733 | -98.6322
21.2771 | -98.6129
21.2808 | -98.5924
21.2846 | -98.5698
21.2883 | -98.547
21.2921 | -98.5248
21.2958 | -98.5028
21.2996 | -98.4771
21.3033 | -98.4516
21.3071 | -98.4244
21.3108 | -98.3973
21.3146 | -98.3725
21.3183 | -98.3486
21.3221 | -98.3255
21.3258 | -98.3002
21.3296 | -98.2755
21.3333 | -98.2514
21.3371 | -98.227
21.3408 | -98.203
21.3446 | -98.1793
21.3483 | -98.1582
21.3521 | -98.1372
21.3558 | -98.1167
21.3596 | -98.0966
21.3633 | -98.0785
21.3671 | -98.0615
21.3708 | -98.0443
21.3746 | -98.0258
21.3783 | -98.01
21.3821 | -97.9958
21.3858 | -97.9847
21.3896 | -97.9738
21.3933 | -97.9606
21.3971 | -97.9487
21.4008 | -97.9374
21.4046 | -97.9262
21.4083 | -97.9158
21.4121 | -97.9085
21.4158 | -97.8991
21.4196 | -97.8904
21.4233 | -97.8843
21.4271 | -97.8783
21.4308 | -97.8742
21.4346 | -97.8715
21.4383 | -97.868
21.4421 | -97.8644
21.4459 | -97.8621
21.4496 | -97.8615
21.4534 | -97.8613
21.4571 | -97.8622
21.4609 | -97.8647
21.4646 | -97.8682
21.4684 | -97.8697
21.4721 | -97.8693
21.4759 | -97.8696
21.4796 | -97.8707
21.4834 | -97.8702
21.4871 | -97.8702
21.4909 | -97.8711
21.4946 | -97.8698
21.4984 | -97.8689
21.5021 | -97.8675
21.5059 | -97.8666
21.5096 | -97.8668
21.5134 | -97.8669
21.5171 | -97.8655
21.5209 | -97.8637
21.5246 | -97.8635
21.5284 | -97.8586
21.5321 | -97.8535
21.5359 | -97.8521
21.5396 | -97.8482
21.5434 | -97.8452
21.5471 | -97.8428
21.5509 | -97.8396
21.5546 | -97.8365
21.5584 | -97.8303
21.5621 | -97.8232
21.5659 | -97.8203
21.5696 | -97.8149
21.5734 | -97.8068
21.5771 | -97.801
21.5809 | -97.7939
21.5846 | -97.7871
21.5884 | -97.7815
21.5921 | -97.7751
21.5959 | -97.7635
21.5996 | -97.751
21.6034 | -97.7425
21.6071 | -97.7323
21.6109 | -97.7249
21.6146 | -97.7183
21.6184 | -97.7111
21.6221 | -97.7044
21.6259 | -97.6978
21.6296 | -97.6919
21.6334 | -97.6863
21.6371 | -97.6818
21.6409 | -97.6809
21.6446 | -97.6775
21.6484 | -97.6747
21.6521 | -97.6696
21.6559 | -97.6645
21.6596 | -97.6636
21.6634 | -97.6637
21.6671 | -97.6624
21.6709 | -97.6591
21.6746 | -97.6557
21.6784 | -97.6545
21.6821 | -97.6575
21.6859 | -97.6589
21.6896 | -97.6607
21.6934 | -97.6629
21.6971 | -97.6671
21.7009 | -97.6735
21.7046 | -97.6791
21.7084 | -97.6844
21.7121 | -97.689
21.7159 | -97.6937
21.7196 | -97.6972
21.7234 | -97.7024
21.7271 | -97.7057
21.7309 | -97.7081
21.7346 | -97.7133
21.7384 | -97.7119
21.7421 | -97.7095
21.7459 | -97.7054
21.7496 | -97.6997
21.7534 | -97.698
21.7571 | -97.6974
21.7609 | -97.6972
21.7647 | -97.6956
21.7684 | -97.6927
21.7722 | -97.6893
21.7759 | -97.6855
21.7797 | -97.6804
21.7834 | -97.6699
21.7872 | -97.6587
21.7909 | -97.6503
21.7947 | -97.638
21.7984 | -97.6219
21.8022 | -97.6062
21.8059 | -97.5886
21.8097 | -97.5701
21.8134 | -97.5509
21.8172 | -97.5304
21.8209 | -97.5061
21.8247 | -97.4817
21.8284 | -97.4573
21.8322 | -97.4324
21.8359 | -97.4086
21.8397 | -97.3839
21.8434 | -97.3542
21.8472 | -97.3246
21.8509 | -97.2977
21.8547 | -97.2718
21.8584 | -97.2474
21.8622 | -97.2226
21.8659 | -97.2007
21.8697 | -97.1787
21.8734 | -97.1557
21.8772 | -97.1368
21.8809 | -97.1187
21.8847 | -97.1008
21.8884 | -97.0863
21.8922 | -97.0743
21.8959 | -97.0649
21.8997 | -97.0586
21.9034 | -97.0579
21.9072 | -97.0578
21.9109 | -97.0589
21.9147 | -97.0641
21.9184 | -97.07
21.9222 | -97.0783
21.9259 | -97.0883
21.9297 | -97.1013
21.9334 | -97.1184
21.9372 | -97.1381
21.9409 | -97.1599
21.9447 | -97.183
21.9484 | -97.2091
21.9522 | -97.2382
21.9559 | -97.2684
21.9597 | -97.3027
21.9634 | -97.3411
21.9672 | -97.3795
21.9709 | -97.4178
21.9747 | -97.4577
21.9784 | -97.5011
21.9822 | -97.5478
21.9859 | -97.5932
21.9897 | -97.639
21.9934 | -97.6852
21.9972 | -97.7309
22.0009 | -97.7757
22.0047 | -97.822
22.0084 | -97.8692
22.0122 | -97.9153
22.0159 | -97.9629
22.0197 | -98.011
22.0234 | -98.059
22.0272 | -98.1069
22.0309 | -98.1551
22.0347 | -98.2004
22.0384 | -98.2455
22.0422 | -98.2896
22.0459 | -98.3295
22.0497 | -98.3732
22.0534 | -98.4176
22.0572 | -98.4605
22.0609 | -98.5001
22.0647 | -98.531
22.0684 | -98.5577
22.0722 | -98.5809
22.0759 | -98.6064
22.0797 | -98.6289
22.0835 | -98.6484
22.0872 | -98.6666
22.091 | -98.6811
22.0947 | -98.6982
22.0985 | -98.7137
22.1022 | -98.7224
22.106 | -98.7312
22.1097 | -98.7398
22.1135 | -98.7432
22.1172 | -98.7451
22.121 | -98.7474
22.1247 | -98.7466
22.1285 | -98.7444
22.1322 | -98.7431
22.136 | -98.7395
22.1397 | -98.7345
22.1435 | -98.7296
22.1472 | -98.7243
22.151 | -98.7171
22.1547 | -98.7117
22.1585 | -98.7077
22.1622 | -98.7021
22.166 | -98.6943
22.1697 | -98.6851
22.1735 | -98.6766
22.1772 | -98.6672
22.181 | -98.6603
22.1847 | -98.6597
22.1885 | -98.6557
22.1922 | -98.6531
22.196 | -98.6509
22.1997 | -98.6484
22.2035 | -98.6463
22.2072 | -98.6457
22.211 | -98.6476
22.2147 | -98.6506
22.2185 | -98.6535
22.2222 | -98.6595
22.226 | -98.6653
22.2297 | -98.6716
22.2335 | -98.679
22.2372 | -98.6862
22.241 | -98.6967
22.2447 | -98.708
22.2485 | -98.7168
22.2522 | -98.7271
22.256 | -98.7385
22.2597 | -98.7491
22.2635 | -98.7609
22.2672 | -98.7758
22.271 | -98.7926
22.2747 | -98.8091
22.2785 | -98.8264
22.2822 | -98.845
22.286 | -98.863
22.2897 | -98.8831
22.2935 | -98.9044
22.2972 | -98.9253
22.301 | -98.9474
22.3047 | -98.9687
22.3085 | -98.9888
22.3122 | -99.0105
22.316 | -99.0304
22.3197 | -99.0499
22.3235 | -99.0683
22.3272 | -99.0842
22.331 | -99.099
22.3347 | -99.1141
22.3385 | -99.1284
22.3422 | -99.1426
22.346 | -99.1564
22.3497 | -99.1722
22.3535 | -99.1869
22.3572 | -99.2013
22.361 | -99.2133
22.3647 | -99.2252
22.3685 | -99.2358
22.3722 | -99.246
22.376 | -99.2573
22.3797 | -99.2679
22.3835 | -99.2767
22.3872 | -99.2829
22.391 | -99.2875
22.3947 | -99.294
22.3985 | -99.3004
22.4022 | -99.3046
22.406 | -99.3105
22.4098 | -99.3162
22.4135 | -99.3186
22.4173 | -99.3218
22.421 | -99.3255
22.4248 | -99.326
22.4285 | -99.3279
22.4323 | -99.3295
22.436 | -99.3304
22.4398 | -99.3331
22.4435 | -99.3363
22.4473 | -99.3385
22.451 | -99.34
22.4548 | -99.3412
22.4585 | -99.3391
22.4623 | -99.3392
22.466 | -99.3407
22.4698 | -99.3423
22.4735 | -99.3449
22.4773 | -99.3449
22.481 | -99.3444
22.4848 | -99.3454
22.4885 | -99.3458
22.4923 | -99.3453
22.496 | -99.3445
22.4998 | -99.3413
22.5035 | -99.336
22.5073 | -99.331
22.511 | -99.3243
22.5148 | -99.3195
22.5185 | -99.3132
22.5223 | -99.3014
22.526 | -99.2869
22.5298 | -99.2735
22.5335 | -99.2618
22.5373 | -99.2519
22.541 | -99.2414
22.5448 | -99.2298
22.5485 | -99.2196
22.5523 | -99.2112
22.556 | -99.2041
22.5598 | -99.1944
22.5635 | -99.1837
22.5673 | -99.1759
22.571 | -99.1673
22.5748 | -99.1604
22.5785 | -99.1549
22.5823 | -99.1475
22.586 | -99.1359
22.5898 | -99.1226
22.5935 | -99.1082
22.5973 | -99.097
22.601 | -99.0879
22.6048 | -99.0776
22.6085 | -99.0675
22.6123 | -99.058
22.616 | -99.0516
22.6198 | -99.0472
22.6235 | -99.0395
22.6273 | -99.0295
22.631 | -99.0227
22.6348 | -99.0211
22.6385 | -99.0174
22.6423 | -99.0117
22.646 | -99.0087
22.6498 | -99.0071
22.6535 | -99.0035
22.6573 | -99.0025
22.661 | -99.0014
22.6648 | -98.9997
22.6685 | -98.9999
22.6723 | -99.0003
22.676 | -98.9988
22.6798 | -99.0013
22.6835 | -99.0042
22.6873 | -99.0047
22.691 | -99.0066
22.6948 | -99.01
22.6985 | -99.0134
22.7023 | -99.0175
22.706 | -99.0262
22.7098 | -99.0324
22.7135 | -99.0388
22.7173 | -99.0463
22.7211 | -99.056
22.7248 | -99.0678
22.7286 | -99.0747
22.7323 | -99.0781
22.7361 | -99.0827
22.7398 | -99.0897
22.7436 | -99.0958
22.7473 | -99.1043
22.7511 | -99.113
22.7548 | -99.1204
22.7586 | -99.1281
22.7623 | -99.1345
22.7661 | -99.1392
22.7698 | -99.1461
22.7736 | -99.1484
22.7773 | -99.1481
22.7811 | -99.1513
22.7848 | -99.1534
22.7886 | -99.157
22.7923 | -99.1607
22.7961 | -99.1618
22.7998 | -99.1615
22.8036 | -99.1634
22.8073 | -99.1607
22.8111 | -99.1566
22.8148 | -99.1538
22.8186 | -99.1502
22.8223 | -99.1454
22.8261 | -99.1387
22.8298 | -99.1302
22.8336 | -99.1178
22.8373 | -99.1088
22.8411 | -99.1017
22.8448 | -99.0974
22.8486 | -99.0919
22.8523 | -99.0791
22.8561 | -99.0624
22.8598 | -99.0471
22.8636 | -99.0318
22.8673 | -99.0139
22.8711 | -98.9945
22.8748 | -98.9765
22.8786 | -98.9599
22.8823 | -98.9446
22.8861 | -98.9298
22.8898 | -98.9147
22.8936 | -98.8998
22.8973 | -98.8851
22.9011 | -98.8712
22.9048 | -98.8579
22.9086 | -98.8471
22.9123 | -98.8365
22.9161 | -98.8252
22.9198 | -98.8112
22.9236 | -98.7974
22.9273 | -98.7902
22.9311 | -98.7834
22.9348 | -98.7772
22.9386 | -98.7714
22.9423 | -98.7682
22.9461 | -98.7688
22.9498 | -98.7713
22.9536 | -98.772
22.9573 | -98.7718
22.9611 | -98.7748
22.9648 | -98.7814
22.9686 | -98.7888
22.9723 | -98.7956
22.9761 | -98.8053
22.9798 | -98.8158
22.9836 | -98.8278
22.9873 | -98.8403
22.9911 | -98.8518
22.9948 | -98.8653
22.9986 | -98.882
23.0023 | -98.8965
23.0061 | -98.9117
23.0098 | -98.9304
23.0136 | -98.9491
23.0173 | -98.9663
23.0211 | -98.9863
23.0248 | -99.0057
23.0286 | -99.021
23.0323 | -99.0398
23.0361 | -99.0604
23.0399 | -99.0783
23.0436 | -99.0908
23.0474 | -99.1021
23.0511 | -99.1115
23.0549 | -99.1248
23.0586 | -99.1406
23.0624 | -99.1506
23.0661 | -99.158
23.0699 | -99.1649
23.0736 | -99.1696
23.0774 | -99.1722
23.0811 | -99.1752
23.0849 | -99.1762
23.0886 | -99.1739
23.0924 | -99.1713
23.0961 | -99.1687
23.0999 | -99.1676
23.1036 | -99.1621
23.1074 | -99.1523
23.1111 | -99.1431
23.1149 | -99.1275
23.1186 | -99.1082
23.1224 | -99.092
23.1261 | -99.0784
23.1299 | -99.0592
23.1336 | -99.0444
23.1374 | -99.0287
23.1411 | -99.0088
23.1449 | -98.9861
23.1486 | -98.9655
23.1524 | -98.9468
23.1561 | -98.9258
23.1599 | -98.907
23.1636 | -98.8851
23.1674 | -98.8619
23.1711 | -98.8413
23.1749 | -98.8206
23.1786 | -98.8022
23.1824 | -98.7856
23.1861 | -98.7678
23.1899 | -98.7518
23.1936 | -98.7392
23.1974 | -98.7289
23.2011 | -98.7172
23.2049 | -98.7049
23.2086 | -98.6959
23.2124 | -98.6883
23.2161 | -98.6829
23.2199 | -98.6799
23.2236 | -98.6768
23.2274 | -98.675
23.2311 | -98.6759
23.2349 | -98.6776
23.2386 | -98.6825
23.2424 | -98.6896
23.2461 | -98.6985
23.2499 | -98.7086
23.2536 | -98.717
23.2574 | -98.7279
23.2611 | -98.7405
23.2649 | -98.764
23.2686 | -98.7777
23.2724 | -98.792
23.2761 | -98.8081
23.2799 | -98.8266
23.2836 | -98.8466
23.2874 | -98.8668
23.2911 | -98.8858
23.2949 | -98.9047
23.2986 | -98.9243
23.3024 | -98.9443
23.3061 | -98.9651
23.3099 | -98.9868
23.3136 | -99.0074
23.3174 | -99.0286
23.3211 | -99.0509
23.3249 | -99.0728
23.3286 | -99.0922
23.3324 | -99.1138
23.3361 | -99.1345
23.3399 | -99.1525
23.3436 | -99.1712
23.3474 | -99.1912
23.3511 | -99.2072
23.3549 | -99.224
23.3587 | -99.2427
23.3624 | -99.2572
23.3662 | -99.2691
23.3699 | -99.2834
23.3737 | -99.3
23.3774 | -99.3158
23.3812 | -99.3307
23.3849 | -99.3423
23.3887 | -99.3491
23.3924 | -99.3573
23.3962 | -99.3642
23.3999 | -99.3693
23.4037 | -99.374
23.4074 | -99.3775
23.4112 | -99.3755
23.4149 | -99.3719
23.4187 | -99.3708
23.4224 | -99.3705
23.4262 | -99.3694
23.4299 | -99.3676
23.4337 | -99.3664
23.4374 | -99.3657
23.4412 | -99.361
23.4449 | -99.3586
23.4487 | -99.3558
23.4524 | -99.3496
23.4562 | -99.3397
23.4599 | -99.335
23.4637 | -99.3273
23.4674 | -99.3202
23.4712 | -99.3159
23.4749 | -99.3123
23.4787 | -99.306
23.4824 | -99.2994
23.4862 | -99.2956
23.4899 | -99.2908
23.4937 | -99.2848
23.4974 | -99.2801
23.5012 | -99.2733
23.5049 | -99.2678
23.5087 | -99.2637
23.5124 | -99.2603
23.5162 | -99.2576
23.5199 | -99.2527
23.5237 | -99.2491
23.5274 | -99.2454
23.5312 | -99.2441
23.5349 | -99.2435
23.5387 | -99.2429
23.5424 | -99.2445
23.5462 | -99.2453
23.5499 | -99.2444
23.5537 | -99.2436
23.5574 | -99.2448
23.5612 | -99.2451
23.5649 | -99.2469
23.5687 | -99.2496
23.5724 | -99.2516
23.5762 | -99.2547
23.5799 | -99.258
23.5837 | -99.2613
23.5874 | -99.2642
23.5912 | -99.2672
23.5949 | -99.271
23.5987 | -99.2749
23.6024 | -99.2794
23.6062 | -99.2823
23.6099 | -99.2864
23.6137 | -99.2909
23.6174 | -99.2943
23.6212 | -99.299
23.6249 | -99.3043
23.6287 | -99.3108
23.6324 | -99.3176
23.6362 | -99.3236
23.6399 | -99.3269
23.6437 | -99.3295
23.6474 | -99.3335
23.6512 | -99.3392
23.6549 | -99.3456
23.6587 | -99.3531
23.6624 | -99.3609
23.6662 | -99.3661
23.6699 | -99.3717
23.6737 | -99.377
23.6775 | -99.3816
23.6812 | -99.3842
23.685 | -99.3851
23.6887 | -99.3846
23.6925 | -99.3846
23.6962 | -99.3864
23.7 | -99.3878
23.7037 | -99.3894
23.7075 | -99.3904
23.7112 | -99.3915
23.715 | -99.3921
23.7187 | -99.3932
23.7225 | -99.3942
23.7262 | -99.3965
23.73 | -99.3959
23.7337 | -99.3945
23.7375 | -99.3935
23.7412 | -99.3915
23.745 | -99.3909
23.7487 | -99.3918
23.7525 | -99.3916
23.7562 | -99.3877
23.76 | -99.3826
23.7637 | -99.3785
23.7675 | -99.3753
23.7712 | -99.3697
23.775 | -99.3646
23.7787 | -99.3598
23.7825 | -99.3575
23.7862 | -99.3539
23.79 | -99.3517
23.7937 | -99.3454
23.7975 | -99.3412
23.8012 | -99.3388
23.805 | -99.3388
23.8087 | -99.3348
23.8125 | -99.3303
23.8162 | -99.3277
23.82 | -99.326
23.8237 | -99.3245
23.8275 | -99.3244
23.8312 | -99.325
23.835 | -99.324
23.8387 | -99.3237
23.8425 | -99.3247
23.8462 | -99.3262
23.85 | -99.328
23.8537 | -99.3297
23.8575 | -99.3333
23.8612 | -99.3363
23.865 | -99.3398
23.8687 | -99.3446
23.8725 | -99.3513
23.8762 | -99.3548
23.88 | -99.3581
23.8837 | -99.3629
23.8875 | -99.3676
23.8912 | -99.3745
23.895 | -99.3806
23.8987 | -99.3859
23.9025 | -99.388
23.9062 | -99.3899
23.91 | -99.3957
23.9137 | -99.401
23.9175 | -99.4069
23.9212 | -99.4116
23.925 | -99.4144
23.9287 | -99.4162
23.9325 | -99.4196
23.9362 | -99.4236
23.94 | -99.4265
23.9437 | -99.4308
23.9475 | -99.4326
23.9512 | -99.4343
23.955 | -99.4374
23.9587 | -99.4406
23.9625 | -99.4429
23.9662 | -99.4432
23.97 | -99.4414
23.9737 | -99.4403
23.9775 | -99.441
23.9812 | -99.441
23.985 | -99.4412
23.9887 | -99.4393
23.9925 | -99.4341
23.9962 | -99.4287
24 | -99.4246
24.0038 | -99.4229
24.0075 | -99.421
24.0113 | -99.4189
24.015 | -99.4167
24.0188 | -99.4153
24.0225 | -99.4173
24.0263 | -99.4189
24.03 | -99.4178
24.0338 | -99.4161
24.0375 | -99.4159
24.0413 | -99.4155
24.045 | -99.4149
24.0488 | -99.4121
24.0525 | -99.4102
24.0563 | -99.4064
24.06 | -99.4021
24.0638 | -99.4004
24.0675 | -99.4005
24.0713 | -99.4005
24.075 | -99.4013
24.0788 | -99.4031
24.0825 | -99.4066
24.0863 | -99.4099
24.09 | -99.4117
24.0938 | -99.4147
24.0975 | -99.4191
24.1013 | -99.4254
24.105 | -99.4316
24.1088 | -99.4376
24.1125 | -99.4431
24.1163 | -99.4475
24.12 | -99.4505
24.1238 | -99.4543
24.1275 | -99.4594
24.1313 | -99.4656
24.135 | -99.4713
24.1388 | -99.4763
24.1425 | -99.4822
24.1463 | -99.4886
24.15 | -99.4941
24.1538 | -99.4972
24.1575 | -99.4994
24.1613 | -99.5018
24.165 | -99.5026
24.1688 | -99.5048
24.1725 | -99.5076
24.1763 | -99.5104
24.18 | -99.5137
24.1838 | -99.5173
24.1875 | -99.5202
24.1913 | -99.5224
24.195 | -99.5258
24.1988 | -99.528
24.2025 | -99.5305
24.2063 | -99.5332
24.21 | -99.5337
24.2138 | -99.5352
24.2175 | -99.539
24.2213 | -99.5424
24.225 | -99.5466
24.2288 | -99.5527
24.2325 | -99.557
24.2363 | -99.56
24.24 | -99.5635
24.2438 | -99.5698
24.2475 | -99.5749
24.2513 | -99.5773
24.255 | -99.5803
24.2588 | -99.5812
24.2625 | -99.5834
24.2663 | -99.5877
24.27 | -99.5917
24.2738 | -99.5927
24.2775 | -99.593
24.2813 | -99.5943
24.285 | -99.5969
24.2888 | -99.5989
24.2925 | -99.6009
24.2963 | -99.6018
24.3 | -99.6024
24.3038 | -99.6024
24.3075 | -99.6016
24.3113 | -99.6023
24.315 | -99.6031
24.3188 | -99.6027
24.3225 | -99.6025
24.3263 | -99.6012
24.3301 | -99.5996
24.3338 | -99.5993
24.3376 | -99.5979
24.3413 | -99.5958
24.3451 | -99.5927
24.3488 | -99.5883
24.3526 | -99.5844
24.3563 | -99.5801
24.3601 | -99.5761
24.3638 | -99.5707
24.3676 | -99.565
24.3713 | -99.5615
24.3751 | -99.5594
24.3788 | -99.5564
24.3826 | -99.5525
24.3863 | -99.5483
24.3901 | -99.5438
24.3938 | -99.54
24.3976 | -99.5388
24.4013 | -99.5402
24.4051 | -99.5429
24.4088 | -99.5391
24.4126 | -99.5363
24.4163 | -99.535
24.4201 | -99.534
24.4238 | -99.5339
24.4276 | -99.5329
24.4313 | -99.5316
24.4351 | -99.5319
24.4388 | -99.5334
24.4426 | -99.5332
24.4463 | -99.5331
24.4501 | -99.5353
24.4538 | -99.5374
24.4576 | -99.5394
24.4613 | -99.5437
24.4651 | -99.5515
24.4688 | -99.5593
24.4726 | -99.5653
24.4763 | -99.5726
24.4801 | -99.5804
24.4838 | -99.5875
24.4876 | -99.5946
24.4913 | -99.6037
24.4951 | -99.615
24.4988 | -99.6291
24.5026 | -99.6416
24.5063 | -99.6525
24.5101 | -99.6662
24.5138 | -99.6794
24.5176 | -99.6923
24.5213 | -99.7069
24.5251 | -99.722
24.5288 | -99.7345
24.5326 | -99.748
24.5363 | -99.7634
24.5401 | -99.7787
24.5438 | -99.7952
24.5476 | -99.8094
24.5513 | -99.8243
24.5551 | -99.8392
24.5588 | -99.8537
24.5626 | -99.8699
24.5663 | -99.8853
24.5701 | -99.9015
24.5738 | -99.9141
24.5776 | -99.9245
24.5813 | -99.9323
24.5851 | -99.9321
24.5888 | -99.9289
24.5926 | -99.9293
24.5963 | -99.9315
24.6001 | -99.9286
24.6038 | -99.9239
24.6076 | -99.9171
24.6113 | -99.9068
24.6151 | -99.8949
24.6188 | -99.8828
24.6226 | -99.8734
24.6263 | -99.8626
24.6301 | -99.847
24.6338 | -99.8281
24.6376 | -99.8065
24.6413 | -99.7827
24.6451 | -99.7532
24.6489 | -99.7211
24.6526 | -99.6881
24.6564 | -99.6529
24.6601 | -99.6146
24.6639 | -99.5732
24.6676 | -99.5304
24.6714 | -99.4876
24.6751 | -99.4425
24.6789 | -99.3945
24.6826 | -99.3459
24.6864 | -99.2966
24.6901 | -99.2478
24.6939 | -99.1991
24.6976 | -99.1493
24.7014 | -99.1009
24.7051 | -99.0542
24.7089 | -99.0069
24.7126 | -98.9614
24.7164 | -98.9175
24.7201 | -98.8742
24.7239 | -98.8313
24.7276 | -98.7899
24.7314 | -98.7516
24.7351 | -98.7166
24.7389 | -98.6868
24.7426 | -98.6608
24.7464 | -98.6354
24.7501 | -98.6124
24.7539 | -98.5929
24.7576 | -98.5798
24.7614 | -98.5683
24.7651 | -98.5597
24.7689 | -98.5551
24.7726 | -98.553
24.7764 | -98.5542
24.7801 | -98.5605
24.7839 | -98.5714
24.7876 | -98.5857
24.7914 | -98.6038
24.7951 | -98.6257
24.7989 | -98.6502
24.8026 | -98.6786
24.8064 | -98.7103
24.8101 | -98.7445
24.8139 | -98.78
24.8176 | -98.8198
24.8214 | -98.863
24.8251 | -98.9088
24.8289 | -98.9574
24.8326 | -99.0089
24.8364 | -99.0621
24.8401 | -99.1158
24.8439 | -99.1707
24.8476 | -99.227
24.8514 | -99.2849
24.8551 | -99.3453
24.8589 | -99.4057
24.8626 | -99.4658
24.8664 | -99.5265
24.8701 | -99.5869
24.8739 | -99.6477
24.8776 | -99.7091
24.8814 | -99.7696
24.8851 | -99.8289
24.8889 | -99.8701
24.8926 | -99.9248
24.8964 | -99.9788
24.9001 | -100.034
24.9039 | -100.0872
24.9076 | -100.1393
24.9114 | -100.1884
24.9151 | -100.2339
24.9189 | -100.2783
24.9226 | -100.3184
24.9264 | -100.3575
24.9301 | -100.3956
24.9339 | -100.4333
24.9376 | -100.4772
24.9414 | -100.5175
24.9451 | -100.5483
24.9489 | -100.573
24.9526 | -100.5982
24.9564 | -100.6225
24.9601 | -100.6435
24.9639 | -100.6627
24.9677 | -100.6777
24.9714 | -100.6884
24.9752 | -100.6934
24.9789 | -100.6957
24.9827 | -100.6982
24.9864 | -100.7025
24.9902 | -100.7047
24.9939 | -100.7066
24.9977 | -100.7102
25.0014 | -100.7146
25.0052 | -100.7178
25.0089 | -100.7196
25.0127 | -100.7197
25.0164 | -100.7204
25.0202 | -100.7201
25.0239 | -100.718
25.0277 | -100.7187
25.0314 | -100.7211
25.0352 | -100.721
25.0389 | -100.7198
25.0427 | -100.7183
25.0464 | -100.7183
25.0502 | -100.7201
25.0539 | -100.7208
25.0577 | -100.7195
25.0614 | -100.7207
25.0652 | -100.7223
25.0689 | -100.7229
25.0727 | -100.7252
25.0764 | -100.7301
25.0802 | -100.7363
25.0839 | -100.7428
25.0877 | -100.7495
25.0914 | -100.7546
25.0952 | -100.7602
25.0989 | -100.7675
25.1027 | -100.7746
25.1064 | -100.7821
25.1102 | -100.7913
25.1139 | -100.8014
25.1177 | -100.8131
25.1214 | -100.8254
25.1252 | -100.8383
25.1289 | -100.8532
25.1327 | -100.8662
25.1364 | -100.8795
25.1402 | -100.8928
25.1439 | -100.9057
25.1477 | -100.918
25.1514 | -100.9326
25.1552 | -100.9481
25.1589 | -100.964
25.1627 | -100.9797
25.1664 | -100.995
25.1702 | -101.009
25.1739 | -101.0241
25.1777 | -101.04
25.1814 | -101.0545
25.1852 | -101.0696
25.1889 | -101.0859
25.1927 | -101.1037
25.1964 | -101.123
25.2002 | -101.1425
25.2039 | -101.1605
25.2077 | -101.1751
25.2114 | -101.1902
25.2152 | -101.2064
25.2189 | -101.222
25.2227 | -101.2356
25.2264 | -101.248
25.2302 | -101.2612
25.2339 | -101.2743
25.2377 | -101.2853
25.2414 | -101.2953
25.2452 | -101.3064
25.2489 | -101.319
25.2527 | -101.3307
25.2564 | -101.3407
25.2602 | -101.3486
25.2639 | -101.3559
25.2677 | -101.3629
25.2714 | -101.3694
25.2752 | -101.3756
25.2789 | -101.3825
25.2827 | -101.3894
25.2865 | -101.3962
25.2902 | -101.4019
25.294 | -101.4047
25.2977 | -101.4075
25.3015 | -101.4086
25.3052 | -101.4084
25.309 | -101.4071
25.3127 | -101.4047
25.3165 | -101.4025
25.3202 | -101.3997
25.324 | -101.3974
25.3277 | -101.3928
25.3315 | -101.3869
25.3352 | -101.3816
25.339 | -101.3762
25.3427 | -101.3704
25.3465 | -101.3655
25.3502 | -101.3584
25.354 | -101.3531
25.3577 | -101.3451
25.3615 | -101.337
25.3652 | -101.3288
25.369 | -101.3212
25.3727 | -101.3126
25.3765 | -101.3034
25.3802 | -101.2947
25.384 | -101.2867
25.3877 | -101.279
25.3915 | -101.2725
25.3952 | -101.2666
25.399 | -101.26
25.4027 | -101.2525
25.4065 | -101.2454
25.4102 | -101.239
25.414 | -101.2335
25.4177 | -101.2261
25.4215 | -101.2194
25.4252 | -101.2136
25.429 | -101.2123
25.4327 | -101.2083
25.4365 | -101.2026
25.4402 | -101.196
25.444 | -101.1895
25.4477 | -101.1824
25.4515 | -101.1768
25.4552 | -101.1723
25.459 | -101.1671
25.4627 | -101.1614
25.4665 | -101.1575
25.4702 | -101.1519
25.474 | -101.1486
25.4777 | -101.144
25.4815 | -101.1381
25.4852 | -101.1354
25.489 | -101.131
25.4927 | -101.1267
25.4965 | -101.1208
25.5002 | -101.1146
25.504 | -101.1112
25.5077 | -101.1085
25.5115 | -101.1037
25.5152 | -101.0987
25.519 | -101.0929
25.5227 | -101.0869
25.5265 | -101.0829
25.5302 | -101.0795
25.534 | -101.0747
25.5377 | -101.0706
25.5415 | -101.067
25.5452 | -101.0628
25.549 | -101.0592
25.5527 | -101.0555
25.5565 | -101.0532
25.5602 | -101.0516
25.564 | -101.0518
25.5677 | -101.0531
25.5715 | -101.0535
25.5752 | -101.053
25.579 | -101.0526
25.5827 | -101.0543
25.5865 | -101.0565
25.5902 | -101.061
25.594 | -101.0655
25.5978 | -101.0714
25.6015 | -101.0791
25.6053 | -101.0868
25.609 | -101.093
25.6128 | -101.0997
25.6165 | -101.1079
25.6203 | -101.1168
25.624 | -101.1267
25.6278 | -101.1371
25.6315 | -101.1502
25.6353 | -101.1633
25.639 | -101.1753
25.6428 | -101.1887
25.6465 | -101.2029
25.6503 | -101.2175
25.654 | -101.2336
25.6578 | -101.2503
25.6615 | -101.2686
25.6653 | -101.2874
25.669 | -101.3071
25.6728 | -101.3276
25.6765 | -101.3486
25.6803 | -101.3704
25.684 | -101.3928
25.6878 | -101.4165
25.6915 | -101.441
25.6953 | -101.4646
25.699 | -101.4876
25.7028 | -101.5111
25.7065 | -101.5352
25.7103 | -101.5621
25.714 | -101.5878
25.7178 | -101.6146
25.7215 | -101.6417
25.7253 | -101.6692
25.729 | -101.6964
25.7328 | -101.7231
25.7365 | -101.7506
25.7403 | -101.7787
25.744 | -101.8044
25.7478 | -101.8286
25.7515 | -101.8523
25.7553 | -101.8751
25.759 | -101.8986
25.7628 | -101.9211
25.7665 | -101.9432
25.7703 | -101.9662
25.774 | -101.9901
25.7778 | -102.0128
25.7815 | -102.0328
25.7853 | -102.0514
25.789 | -102.068
25.7928 | -102.081
25.7965 | -102.0951
25.8003 | -102.1105
25.804 | -102.1261
25.8078 | -102.1396
25.8115 | -102.1529
25.8153 | -102.1651
25.819 | -102.1743
25.8228 | -102.1837
25.8265 | -102.1954
25.8303 | -102.205
25.834 | -102.2144
25.8378 | -102.2199
25.8415 | -102.2254
25.8453 | -102.2322
25.849 | -102.2361
25.8528 | -102.2396
25.8565 | -102.2424
25.8603 | -102.2468
25.864 | -102.2515
25.8678 | -102.2528
25.8715 | -102.2534
25.8753 | -102.2568
25.879 | -102.2585
25.8828 | -102.2601
25.8865 | -102.2619
25.8903 | -102.2604
25.894 | -102.2584
25.8978 | -102.2577
25.9015 | -102.2569
25.9053 | -102.2565
25.909 | -102.2566
25.9128 | -102.256
25.9165 | -102.2558
25.9203 | -102.2538
25.9241 | -102.251
25.9278 | -102.2501
25.9316 | -102.2507
25.9353 | -102.2516
25.9391 | -102.2539
25.9428 | -102.2558
25.9466 | -102.2571
25.9503 | -102.2576
25.9541 | -102.2582
25.9578 | -102.258
25.9616 | -102.2577
25.9653 | -102.2584
25.9691 | -102.2606
25.9728 | -102.2612
25.9766 | -102.263
25.9803 | -102.2649
25.9841 | -102.2671
25.9878 | -102.2705
25.9916 | -102.2728
25.9953 | -102.2761
25.9991 | -102.2788
26.0028 | -102.2811
26.0066 | -102.2843
26.0103 | -102.2877
26.0141 | -102.2896
26.0178 | -102.2925
26.0216 | -102.2954
26.0253 | -102.2972
26.0291 | -102.3008
26.0328 | -102.3072
26.0366 | -102.3131
26.0403 | -102.3185
26.0441 | -102.3239
26.0478 | -102.328
26.0516 | -102.3336
26.0553 | -102.339
26.0591 | -102.3426
26.0628 | -102.3482
26.0666 | -102.3563
26.0703 | -102.3644
26.0741 | -102.3707
26.0778 | -102.3764
26.0816 | -102.3832
26.0853 | -102.3901
26.0891 | -102.3959
26.0928 | -102.4009
26.0966 | -102.4061
26.1003 | -102.4132
26.1041 | -102.4203
26.1078 | -102.4262
26.1116 | -102.4324
26.1153 | -102.4399
26.1191 | -102.4478
26.1228 | -102.4556
26.1266 | -102.4641
26.1303 | -102.4726
26.1341 | -102.4804
26.1378 | -102.4878
26.1416 | -102.4944
26.1453 | -102.502
26.1491 | -102.5117
26.1528 | -102.5217
26.1566 | -102.5305
26.1603 | -102.5402
26.1641 | -102.5476
26.1678 | -102.5541
26.1716 | -102.5606
26.1753 | -102.5693
26.1791 | -102.5769
26.1828 | -102.5833
26.1866 | -102.5917
26.1903 | -102.6005
26.1941 | -102.6064
26.1978 | -102.6119
26.2016 | -102.6196
26.2053 | -102.6286
26.2091 | -102.6374
26.2128 | -102.6457
26.2166 | -102.6533
26.2203 | -102.6608
26.2241 | -102.6688
26.2278 | -102.6752
26.2316 | -102.6811
26.2353 | -102.6868
26.2391 | -102.6931
26.2429 | -102.6964
26.2466 | -102.7011
26.2504 | -102.7068
26.2541 | -102.7107
26.2579 | -102.7162
26.2616 | -102.7211
26.2654 | -102.7255
26.2691 | -102.7297
26.2729 | -102.7337
26.2766 | -102.7365
26.2804 | -102.7402
26.2841 | -102.7448
26.2879 | -102.7487
26.2916 | -102.7538
26.2954 | -102.7581
26.2991 | -102.7622
26.3029 | -102.7646
26.3066 | -102.7664
26.3104 | -102.7689
26.3141 | -102.7721
26.3179 | -102.7762
26.3216 | -102.7788
26.3254 | -102.7819
26.3291 | -102.7853
26.3329 | -102.7884
26.3366 | -102.7931
26.3404 | -102.7989
26.3441 | -102.8047
26.3479 | -102.8107
26.3516 | -102.8151
26.3554 | -102.8178
26.3591 | -102.8207
26.3629 | -102.8252
26.3666 | -102.8294
26.3704 | -102.8326
26.3741 | -102.8354
26.3779 | -102.8384
26.3816 | -102.8427
26.3854 | -102.8461
26.3891 | -102.8473
26.3929 | -102.8487
26.3966 | -102.8527
26.4004 | -102.8542
26.4041 | -102.8531
26.4079 | -102.8512
26.4116 | -102.8482
26.4154 | -102.8473
26.4191 | -102.8484
26.4229 | -102.85
26.4266 | -102.8505
26.4304 | -102.8496
26.4341 | -102.8482
26.4379 | -102.8467
26.4416 | -102.8443
26.4454 | -102.841
26.4491 | -102.8394
26.4529 | -102.8367
26.4566 | -102.833
26.4604 | -102.8294
26.4641 | -102.8261
26.4679 | -102.8229
26.4716 | -102.8198
26.4754 | -102.8168
26.4791 | -102.8122
26.4829 | -102.8088
26.4866 | -102.8061
26.4904 | -102.8031
26.4941 | -102.8009
26.4979 | -102.7933
26.5016 | -102.7873
26.5054 | -102.7813
26.5091 | -102.7758
26.5129 | -102.7718
26.5166 | -102.7677
26.5204 | -102.7635
26.5241 | -102.7579
26.5279 | -102.7522
26.5316 | -102.7478
26.5354 | -102.7436
26.5391 | -102.7394
26.5429 | -102.7343
26.5466 | -102.7293
26.5504 | -102.7242
26.5541 | -102.7202
26.5579 | -102.7152
26.5617 | -102.7096
26.5654 | -102.7031
26.5692 | -102.6984
26.5729 | -102.6941
26.5767 | -102.6887
26.5804 | -102.6821
26.5842 | -102.6748
26.5879 | -102.6672
26.5917 | -102.6617
26.5954 | -102.6584
26.5992 | -102.6536
26.6029 | -102.6471
26.6067 | -102.6408
26.6104 | -102.6375
26.6142 | -102.6336
26.6179 | -102.6298
26.6217 | -102.6262
26.6254 | -102.6221
26.6292 | -102.6179
26.6329 | -102.6126
26.6367 | -102.6065
26.6404 | -102.6004
26.6442 | -102.5952
26.6479 | -102.5915
26.6517 | -102.5894
26.6554 | -102.5846
26.6592 | -102.579
26.6629 | -102.5755
26.6667 | -102.571
26.6704 | -102.567
26.6742 | -102.5636
26.6779 | -102.5624
26.6817 | -102.5612
26.6854 | -102.559
26.6892 | -102.5546
26.6929 | -102.5497
26.6967 | -102.5459
26.7004 | -102.543
26.7042 | -102.5395
26.7079 | -102.535
26.7117 | -102.5295
26.7154 | -102.5223
26.7192 | -102.5144
26.7229 | -102.5063
26.7267 | -102.5
26.7304 | -102.4964
26.7342 | -102.4905
26.7379 | -102.4858
26.7417 | -102.48
26.7454 | -102.4743
26.7492 | -102.4689
26.7529 | -102.462
26.7567 | -102.4536
26.7604 | -102.4448
26.7642 | -102.4365
26.7679 | -102.4289
26.7717 | -102.4195
26.7754 | -102.4107
26.7792 | -102.4021
26.7829 | -102.3947
26.7867 | -102.3884
26.7904 | -102.3794
26.7942 | -102.3703
26.7979 | -102.3601
26.8017 | -102.3505
26.8054 | -102.3427
26.8092 | -102.3347
26.8129 | -102.3262
26.8167 | -102.3184
26.8204 | -102.3103
26.8242 | -102.2992
26.8279 | -102.2918
26.8317 | -102.2859
26.8354 | -102.2803
26.8392 | -102.2747
26.8429 | -102.2681
26.8467 | -102.2611
26.8504 | -102.2554
26.8542 | -102.2516
26.8579 | -102.2469
26.8617 | -102.2427
26.8654 | -102.2399
26.8692 | -102.2363
26.8729 | -102.2362
26.8767 | -102.2346
26.8804 | -102.2339
26.8842 | -102.2334
26.888 | -102.2338
26.8917 | -102.237
26.8955 | -102.2409
26.8992 | -102.2456
26.903 | -102.2508
26.9067 | -102.2575
26.9105 | -102.2636
26.9142 | -102.2715
26.918 | -102.2804
26.9217 | -102.2899
26.9255 | -102.3001
26.9292 | -102.3098
26.933 | -102.3203
26.9367 | -102.3325
26.9405 | -102.345
26.9442 | -102.3593
26.948 | -102.3746
26.9517 | -102.3908
26.9555 | -102.4077
26.9592 | -102.4262
26.963 | -102.445
26.9667 | -102.4653
26.9705 | -102.485
26.9742 | -102.5022
26.978 | -102.5212
26.9817 | -102.5427
26.9855 | -102.5634
26.9892 | -102.5808
26.993 | -102.599
26.9967 | -102.6177
27.0005 | -102.6359
27.0042 | -102.6541
27.008 | -102.6736
27.0117 | -102.6937
27.0155 | -102.7153
27.0192 | -102.736
27.023 | -102.7542
27.0267 | -102.7708
27.0305 | -102.7873
27.0342 | -102.8068
27.038 | -102.8261
27.0417 | -102.8427
27.0455 | -102.8594
27.0492 | -102.8753
27.053 | -102.8853
27.0567 | -102.8986
27.0605 | -102.9111
27.0642 | -102.9245
27.068 | -102.9385
27.0717 | -102.9522
27.0755 | -102.9656
27.0792 | -102.9764
27.083 | -102.9862
27.0867 | -102.9957
27.0905 | -103.0049
27.0942 | -103.0155
27.098 | -103.0251
27.1017 | -103.0333
27.1055 | -103.0409
27.1092 | -103.0491
27.113 | -103.0542
27.1167 | -103.0584
27.1205 | -103.0609
27.1242 | -103.0638
27.128 | -103.0677
27.1317 | -103.071
27.1355 | -103.0737
27.1392 | -103.077
27.143 | -103.0824
27.1467 | -103.0862
27.1505 | -103.0875
27.1542 | -103.0884
27.158 | -103.0888
27.1617 | -103.0898
27.1655 | -103.0907
27.1692 | -103.0913
27.173 | -103.0912
27.1767 | -103.0885
27.1805 | -103.0874
27.1842 | -103.0862
27.188 | -103.083
27.1917 | -103.0814
27.1955 | -103.0802
27.1992 | -103.0788
27.203 | -103.0778
27.2068 | -103.0788
27.2105 | -103.0763
27.2143 | -103.0727
27.218 | -103.0681
27.2218 | -103.0645
27.2255 | -103.0618
27.2293 | -103.0589
27.233 | -103.055
27.2368 | -103.0516
27.2405 | -103.0485
27.2443 | -103.0444
27.248 | -103.0399
27.2518 | -103.0357
27.2555 | -103.0316
27.2593 | -103.0274
27.263 | -103.0229
27.2668 | -103.0194
27.2705 | -103.0155
27.2743 | -103.0109
27.278 | -103.0061
27.2818 | -103.003
27.2855 | -103.0017
27.2893 | -102.9991
27.293 | -102.998
27.2968 | -102.997
27.3005 | -102.996
27.3043 | -102.9959
27.308 | -102.996
27.3118 | -102.9973
27.3155 | -102.9988
27.3193 | -103.0008
27.323 | -103.0037
27.3268 | -103.0059
27.3305 | -103.0075
27.3343 | -103.009
27.338 | -103.0132
27.3418 | -103.0168
27.3455 | -103.0227
27.3493 | -103.0279
27.353 | -103.0337
27.3568 | -103.0406
27.3605 | -103.048
27.3643 | -103.0552
27.368 | -103.0623
27.3718 | -103.0702
27.3755 | -103.0781
27.3793 | -103.0867
27.383 | -103.0954
27.3868 | -103.1052
27.3905 | -103.1147
27.3943 | -103.1234
27.398 | -103.1317
27.4018 | -103.1398
27.4055 | -103.1463
27.4093 | -103.1544
27.413 | -103.1632
27.4168 | -103.1706
27.4205 | -103.1772
27.4243 | -103.1847
27.428 | -103.1923
27.4318 | -103.2006
27.4355 | -103.2087
27.4393 | -103.2145
27.443 | -103.2205
27.4468 | -103.2254
27.4505 | -103.2286
27.4543 | -103.2343
27.458 | -103.2399
27.4618 | -103.2443
27.4655 | -103.2492
27.4693 | -103.2524
27.473 | -103.2567
27.4768 | -103.2607
27.4805 | -103.262
27.4843 | -103.2635
27.488 | -103.2655
27.4918 | -103.2676
27.4955 | -103.2672
27.4993 | -103.2683
27.503 | -103.2696
27.5068 | -103.2689
27.5105 | -103.267
27.5143 | -103.261
27.518 | -103.2547
27.5218 | -103.251
27.5256 | -103.2464
27.5293 | -103.2413
27.5331 | -103.2365
27.5368 | -103.232
27.5406 | -103.2264
27.5443 | -103.2207
27.5481 | -103.2151
27.5518 | -103.209
27.5556 | -103.2021
27.5593 | -103.1957
27.5631 | -103.1909
27.5668 | -103.1842
27.5706 | -103.1783
27.5743 | -103.1739
27.5781 | -103.1675
27.5818 | -103.1608
27.5856 | -103.1551
27.5893 | -103.1508
27.5931 | -103.1488
27.5968 | -103.1437
27.6006 | -103.1384
27.6043 | -103.1339
27.6081 | -103.1308
27.6118 | -103.1278
27.6156 | -103.1247
27.6193 | -103.1213
27.6231 | -103.1182
27.6268 | -103.116
27.6306 | -103.1156
27.6343 | -103.1161
27.6381 | -103.1163
27.6418 | -103.1181
27.6456 | -103.12
27.6493 | -103.1223
27.6531 | -103.1242
27.6568 | -103.1258
27.6606 | -103.1272
27.6643 | -103.1289
27.6681 | -103.1317
27.6718 | -103.1332
27.6756 | -103.1352
27.6793 | -103.1386
27.6831 | -103.1404
27.6868 | -103.143
27.6906 | -103.147
27.6943 | -103.1516
27.6981 | -103.1547
27.7018 | -103.1585
27.7056 | -103.163
27.7093 | -103.1675
27.7131 | -103.1722
27.7168 | -103.1769
27.7206 | -103.1815
27.7243 | -103.1865
27.7281 | -103.1898
27.7318 | -103.195
27.7356 | -103.1984
27.7393 | -103.2013
27.7431 | -103.2044
27.7468 | -103.208
27.7506 | -103.212
27.7543 | -103.2146
27.7581 | -103.2167
27.7618 | -103.219
27.7656 | -103.2214
27.7693 | -103.2247
27.7731 | -103.2273
27.7768 | -103.2288
27.7806 | -103.229
27.7843 | -103.229
27.7881 | -103.2292
27.7918 | -103.2303
27.7956 | -103.2335
27.7993 | -103.234
27.8031 | -103.2317
27.8068 | -103.2311
27.8106 | -103.2323
27.8143 | -103.2332
27.8181 | -103.2331
27.8218 | -103.232
27.8256 | -103.233
27.8293 | -103.2357
27.8331 | -103.2358
27.8368 | -103.2362
27.8406 | -103.2363
27.8444 | -103.234
27.8481 | -103.2306
27.8519 | -103.2261
27.8556 | -103.2237
27.8594 | -103.2216
27.8631 | -103.2175
27.8669 | -103.2128
27.8706 | -103.2091
27.8744 | -103.2057
27.8781 | -103.2034
27.8819 | -103.2014
27.8856 | -103.1999
27.8894 | -103.1964
27.8931 | -103.1938
27.8969 | -103.1917
27.9006 | -103.1893
27.9044 | -103.1888
27.9081 | -103.1873
27.9119 | -103.1844
27.9156 | -103.1805
27.9194 | -103.1753
27.9231 | -103.1723
27.9269 | -103.1688
27.9306 | -103.1653
27.9344 | -103.1605
27.9381 | -103.1552
27.9419 | -103.1515
27.9456 | -103.1472
27.9494 | -103.1432
27.9531 | -103.1401
27.9569 | -103.1366
27.9606 | -103.1327
27.9644 | -103.1295
27.9681 | -103.1258
27.9719 | -103.1219
27.9756 | -103.1196
27.9794 | -103.1182
27.9831 | -103.1149
27.9869 | -103.1133
27.9906 | -103.1114
27.9944 | -103.11
27.9981 | -103.1084
28.0019 | -103.1061