## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 243 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 260
Hood | 02-92B/P | 187
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1005 | hPa |  | PRC PRO 01
Humidity | 54.4 | %/Hr
Internal temperature of calibration bench | 23.97 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1013 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.3 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.98 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3112 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 06/06/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 11/06/2018
Technical Manager | Anthony HUE | Approval |  | 13/06/2018
CEO | Robert GLEMOT | Approval |  | 13/06/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.495 | 57.5359
0.012209 | -36.0557 | 38.3406
0.018313 | -34.9837 | 28.2281
0.024417 | -34.5193 | 22.0017
0.030521 | -34.4123 | 18.2503
0.036626 | -34.267 | 15.1951
0.04273 | -34.1704 | 12.8125
0.048834 | -34.1246 | 11.3032
0.054938 | -34.0566 | 10.1783
0.061043 | -33.9928 | 9.0895
0.067147 | -33.9828 | 8.2693
0.073251 | -33.9848 | 7.5129
0.079355 | -34.0053 | 6.8543
0.08546 | -34.0207 | 6.2584
0.091564 | -34.0153 | 5.8191
0.097668 | -33.9952 | 5.5813
0.10377 | -33.9709 | 5.3416
0.10988 | -33.952 | 5.0685
0.11598 | -33.942 | 4.7759
0.12209 | -33.9439 | 4.4904
0.12819 | -33.9547 | 4.2029
0.13429 | -33.9665 | 3.9089
0.1404 | -33.9703 | 3.6236
0.1465 | -33.9725 | 3.3604
0.15261 | -33.9697 | 3.1368
0.15871 | -33.9628 | 2.9747
0.16482 | -33.954 | 2.8757
0.17092 | -33.9512 | 2.8432
0.17702 | -33.9485 | 2.7752
0.18313 | -33.9472 | 2.7
0.18923 | -33.9465 | 2.5913
0.19534 | -33.9448 | 2.4633
0.20144 | -33.9431 | 2.3186
0.20754 | -33.9398 | 2.1967
0.21365 | -33.9385 | 2.0813
0.21975 | -33.9378 | 1.9835
0.22586 | -33.9388 | 1.8966
0.23196 | -33.9413 | 1.8253
0.23807 | -33.9455 | 1.7641
0.24417 | -33.9497 | 1.7069
0.25027 | -33.9538 | 1.6457
0.25638 | -33.9562 | 1.5849
0.26248 | -33.958 | 1.5188
0.26859 | -33.9589 | 1.4539
0.27469 | -33.9591 | 1.3826
0.2808 | -33.96 | 1.3078
0.2869 | -33.9607 | 1.236
0.293 | -33.9615 | 1.1681
0.29911 | -33.9624 | 1.1075
0.30521 | -33.9633 | 1.0478
0.31132 | -33.9639 | 1.0095
0.31742 | -33.9651 | 0.96681
0.32353 | -33.9653 | 0.93256
0.32963 | -33.9661 | 0.90458
0.33573 | -33.9671 | 0.88608
0.34184 | -33.9684 | 0.86639
0.34794 | -33.9705 | 0.84432
0.35405 | -33.9727 | 0.82184
0.36015 | -33.9753 | 0.79803
0.36626 | -33.9773 | 0.77191
0.37236 | -33.9795 | 0.73583
0.37846 | -33.9817 | 0.69151
0.38457 | -33.9825 | 0.63287
0.39067 | -33.9835 | 0.58561
0.39678 | -33.9841 | 0.53259
0.40288 | -33.9844 | 0.47712
0.40899 | -33.9842 | 0.42381
0.41509 | -33.9834 | 0.3715
0.42119 | -33.9832 | 0.32724
0.4273 | -33.983 | 0.2799
0.4334 | -33.9828 | 0.23264
0.43951 | -33.9829 | 0.18837
0.44561 | -33.9837 | 0.14798
0.45172 | -33.9852 | 0.11621
0.45782 | -33.986 | 0.078619
0.46392 | -33.9872 | 0.042
0.47003 | -33.9882 | 0.0094192
0.47613 | -33.989 | -0.025026
0.48224 | -33.9894 | -0.060315
0.48834 | -33.9884 | -0.095707
0.49445 | -33.988 | -0.13161
0.50055 | -33.9873 | -0.16406
0.50665 | -33.9865 | -0.19886
0.51276 | -33.9853 | -0.23355
0.51886 | -33.9841 | -0.26817
0.52497 | -33.9827 | -0.29236
0.53107 | -33.9817 | -0.32219
0.53717 | -33.9809 | -0.35069
0.54328 | -33.9799 | -0.37341
0.54938 | -33.9796 | -0.39581
0.55549 | -33.9795 | -0.41652
0.56159 | -33.9806 | -0.42735
0.5677 | -33.9808 | -0.44136
0.5738 | -33.9818 | -0.45722
0.5799 | -33.9826 | -0.47167
0.58601 | -33.9837 | -0.48101
0.59211 | -33.9854 | -0.48633
0.59822 | -33.9868 | -0.49502
0.60432 | -33.9884 | -0.50649
0.61043 | -33.9897 | -0.51357
0.61653 | -33.9909 | -0.52318
0.62263 | -33.9925 | -0.53183
0.62874 | -33.9932 | -0.5419
0.63484 | -33.9941 | -0.55046
0.64095 | -33.9945 | -0.56003
0.64705 | -33.9949 | -0.56911
0.65316 | -33.995 | -0.58109
0.65926 | -33.9951 | -0.59667
0.66536 | -33.9938 | -0.61395
0.67147 | -33.9933 | -0.62671
0.67757 | -33.9929 | -0.64067
0.68368 | -33.992 | -0.65953
0.68978 | -33.9911 | -0.67659
0.69589 | -33.99 | -0.69557
0.70199 | -33.9889 | -0.71861
0.70809 | -33.9878 | -0.73783
0.7142 | -33.987 | -0.75844
0.7203 | -33.9857 | -0.77768
0.72641 | -33.9849 | -0.80047
0.73251 | -33.9841 | -0.82061
0.73862 | -33.9843 | -0.84273
0.74472 | -33.9839 | -0.8629
0.75082 | -33.9835 | -0.88531
0.75693 | -33.9838 | -0.90745
0.76303 | -33.9843 | -0.92979
0.76914 | -33.9862 | -0.95055
0.77524 | -33.9871 | -0.97039
0.78135 | -33.9884 | -0.99068
0.78745 | -33.9898 | -1.0109
0.79355 | -33.991 | -1.0308
0.79966 | -33.9928 | -1.0505
0.80576 | -33.9952 | -1.0704
0.81187 | -33.9969 | -1.0923
0.81797 | -33.9989 | -1.1138
0.82408 | -34.0007 | -1.1346
0.83018 | -34.0025 | -1.1555
0.83628 | -34.0044 | -1.174
0.84239 | -34.0056 | -1.1949
0.84849 | -34.0071 | -1.2169
0.8546 | -34.0086 | -1.237
0.8607 | -34.0099 | -1.258
0.86681 | -34.011 | -1.2771
0.87291 | -34.012 | -1.2961
0.87901 | -34.0125 | -1.3143
0.88512 | -34.0132 | -1.3326
0.89122 | -34.0139 | -1.351
0.89733 | -34.0143 | -1.3666
0.90343 | -34.0147 | -1.3844
0.90953 | -34.0148 | -1.3987
0.91564 | -34.0148 | -1.4146
0.92174 | -34.015 | -1.4295
0.92785 | -34.015 | -1.4441
0.93395 | -34.0151 | -1.4579
0.94006 | -34.0152 | -1.4706
0.94616 | -34.0148 | -1.4823
0.95226 | -34.0149 | -1.4948
0.95837 | -34.0149 | -1.5086
0.96447 | -34.0151 | -1.5213
0.97058 | -34.0148 | -1.5346
0.97668 | -34.0145 | -1.5478
0.98279 | -34.0142 | -1.5604
0.98889 | -34.0139 | -1.5708
0.99499 | -34.0138 | -1.5841
1.0011 | -34.0137 | -1.5945
1.0072 | -34.0132 | -1.6058
1.0133 | -34.0128 | -1.6163
1.0194 | -34.0123 | -1.6291
1.0255 | -34.0119 | -1.6427
1.0316 | -34.0112 | -1.6544
1.0377 | -34.0108 | -1.6679
1.0438 | -34.0101 | -1.6805
1.0499 | -34.0092 | -1.6957
1.056 | -34.0084 | -1.71
1.0621 | -34.0077 | -1.7242
1.0682 | -34.0071 | -1.7364
1.0743 | -34.0064 | -1.7519
1.0805 | -34.0058 | -1.7662
1.0866 | -34.0047 | -1.7812
1.0927 | -34.0038 | -1.7956
1.0988 | -34.003 | -1.8105
1.1049 | -34.0021 | -1.8246
1.111 | -34.0012 | -1.84
1.1171 | -34.0004 | -1.8544
1.1232 | -33.9998 | -1.8696
1.1293 | -33.9993 | -1.8835
1.1354 | -33.999 | -1.8985
1.1415 | -33.9985 | -1.914
1.1476 | -33.9978 | -1.9295
1.1537 | -33.9972 | -1.9442
1.1598 | -33.9972 | -1.9589
1.1659 | -33.9969 | -1.972
1.172 | -33.9966 | -1.9858
1.1781 | -33.9963 | -1.9997
1.1842 | -33.9961 | -2.0137
1.1903 | -33.9965 | -2.0244
1.1964 | -33.9963 | -2.0377
1.2025 | -33.9964 | -2.0515
1.2086 | -33.9967 | -2.0636
1.2147 | -33.9967 | -2.0759
1.2209 | -33.997 | -2.0868
1.227 | -33.9977 | -2.0972
1.2331 | -33.9981 | -2.109
1.2392 | -33.9985 | -2.1205
1.2453 | -33.9989 | -2.1313
1.2514 | -33.9993 | -2.1408
1.2575 | -33.9999 | -2.1515
1.2636 | -34.0007 | -2.16
1.2697 | -34.0014 | -2.1688
1.2758 | -34.002 | -2.1797
1.2819 | -34.0028 | -2.1893
1.288 | -34.0034 | -2.1997
1.2941 | -34.0043 | -2.2078
1.3002 | -34.005 | -2.2154
1.3063 | -34.0055 | -2.224
1.3124 | -34.0062 | -2.2326
1.3185 | -34.0069 | -2.2413
1.3246 | -34.0076 | -2.25
1.3307 | -34.0085 | -2.2582
1.3368 | -34.0093 | -2.2665
1.3429 | -34.0099 | -2.2751
1.349 | -34.0104 | -2.284
1.3551 | -34.0111 | -2.293
1.3613 | -34.0117 | -2.3009
1.3674 | -34.0123 | -2.3107
1.3735 | -34.0127 | -2.3201
1.3796 | -34.0132 | -2.3288
1.3857 | -34.0135 | -2.3377
1.3918 | -34.014 | -2.3476
1.3979 | -34.0146 | -2.3564
1.404 | -34.015 | -2.3675
1.4101 | -34.0154 | -2.3776
1.4162 | -34.0158 | -2.3879
1.4223 | -34.0163 | -2.3991
1.4284 | -34.0167 | -2.4088
1.4345 | -34.017 | -2.4201
1.4406 | -34.0172 | -2.4335
1.4467 | -34.0177 | -2.4449
1.4528 | -34.0181 | -2.456
1.4589 | -34.0186 | -2.4676
1.465 | -34.0188 | -2.4782
1.4711 | -34.0193 | -2.4893
1.4772 | -34.0198 | -2.5025
1.4833 | -34.0202 | -2.5154
1.4894 | -34.0208 | -2.5275
1.4955 | -34.0213 | -2.5392
1.5016 | -34.0218 | -2.5516
1.5078 | -34.0222 | -2.5654
1.5139 | -34.0226 | -2.5783
1.52 | -34.023 | -2.5907
1.5261 | -34.0233 | -2.6031
1.5322 | -34.0236 | -2.6171
1.5383 | -34.024 | -2.6295
1.5444 | -34.0245 | -2.6419
1.5505 | -34.025 | -2.6552
1.5566 | -34.0255 | -2.6681
1.5627 | -34.0259 | -2.6818
1.5688 | -34.0265 | -2.6953
1.5749 | -34.0269 | -2.7077
1.581 | -34.0271 | -2.7209
1.5871 | -34.0275 | -2.7346
1.5932 | -34.0279 | -2.7473
1.5993 | -34.0283 | -2.7614
1.6054 | -34.0286 | -2.7745
1.6115 | -34.029 | -2.7886
1.6176 | -34.0293 | -2.801
1.6237 | -34.0295 | -2.8145
1.6298 | -34.0297 | -2.8269
1.6359 | -34.0301 | -2.8395
1.642 | -34.0305 | -2.8512
1.6482 | -34.0306 | -2.8641
1.6543 | -34.0309 | -2.8776
1.6604 | -34.0313 | -2.8895
1.6665 | -34.0313 | -2.9017
1.6726 | -34.0315 | -2.9141
1.6787 | -34.0318 | -2.9265
1.6848 | -34.0317 | -2.9389
1.6909 | -34.0319 | -2.9515
1.697 | -34.0321 | -2.9635
1.7031 | -34.0321 | -2.9759
1.7092 | -34.032 | -2.9879
1.7153 | -34.0321 | -2.9997
1.7214 | -34.0319 | -3.0118
1.7275 | -34.0317 | -3.0243
1.7336 | -34.0316 | -3.0371
1.7397 | -34.0317 | -3.0493
1.7458 | -34.0317 | -3.0614
1.7519 | -34.0316 | -3.0734
1.758 | -34.0316 | -3.0852
1.7641 | -34.0317 | -3.0967
1.7702 | -34.0316 | -3.1095
1.7763 | -34.0315 | -3.1208
1.7824 | -34.0315 | -3.1331
1.7885 | -34.0314 | -3.1456
1.7947 | -34.0315 | -3.1575
1.8008 | -34.0314 | -3.1685
1.8069 | -34.0314 | -3.18
1.813 | -34.0314 | -3.192
1.8191 | -34.0313 | -3.2047
1.8252 | -34.0314 | -3.2175
1.8313 | -34.0315 | -3.2296
1.8374 | -34.0316 | -3.2414
1.8435 | -34.0316 | -3.2532
1.8496 | -34.0316 | -3.2663
1.8557 | -34.0318 | -3.279
1.8618 | -34.032 | -3.2913
1.8679 | -34.0321 | -3.3032
1.874 | -34.0321 | -3.3151
1.8801 | -34.0323 | -3.3275
1.8862 | -34.0325 | -3.3394
1.8923 | -34.0326 | -3.3515
1.8984 | -34.0327 | -3.3636
1.9045 | -34.0327 | -3.3751
1.9106 | -34.0327 | -3.3873
1.9167 | -34.0327 | -3.3995
1.9228 | -34.0329 | -3.4113
1.9289 | -34.0329 | -3.4231
1.9351 | -34.0331 | -3.4339
1.9412 | -34.0333 | -3.445
1.9473 | -34.0336 | -3.4571
1.9534 | -34.0337 | -3.4684
1.9595 | -34.0339 | -3.4802
1.9656 | -34.0342 | -3.4925
1.9717 | -34.0344 | -3.5042
1.9778 | -34.0347 | -3.5155
1.9839 | -34.0349 | -3.527
1.99 | -34.0351 | -3.538
1.9961 | -34.0351 | -3.5501
2.0022 | -34.0354 | -3.5615
2.0083 | -34.0354 | -3.5726
2.0144 | -34.0356 | -3.584
2.0205 | -34.0359 | -3.5957
2.0266 | -34.0363 | -3.6071
2.0327 | -34.0365 | -3.6186
2.0388 | -34.0369 | -3.6301
2.0449 | -34.0371 | -3.6415
2.051 | -34.0373 | -3.6528
2.0571 | -34.0377 | -3.6641
2.0632 | -34.0378 | -3.6755
2.0693 | -34.038 | -3.6877
2.0754 | -34.0383 | -3.699
2.0816 | -34.0386 | -3.7112
2.0877 | -34.0388 | -3.7226
2.0938 | -34.0391 | -3.7347
2.0999 | -34.0393 | -3.7464
2.106 | -34.0395 | -3.758
2.1121 | -34.0397 | -3.77
2.1182 | -34.0399 | -3.7806
2.1243 | -34.0402 | -3.7913
2.1304 | -34.0404 | -3.8024
2.1365 | -34.0407 | -3.8141
2.1426 | -34.0409 | -3.826
2.1487 | -34.0412 | -3.8372
2.1548 | -34.0416 | -3.849
2.1609 | -34.042 | -3.8604
2.167 | -34.0424 | -3.8715
2.1731 | -34.0427 | -3.8835
2.1792 | -34.0429 | -3.8945
2.1853 | -34.0433 | -3.9064
2.1914 | -34.0436 | -3.9178
2.1975 | -34.044 | -3.9293
2.2036 | -34.0442 | -3.9414
2.2097 | -34.0444 | -3.9528
2.2158 | -34.0448 | -3.9633
2.222 | -34.045 | -3.9737
2.2281 | -34.0453 | -3.985
2.2342 | -34.0456 | -3.9963
2.2403 | -34.0457 | -4.0068
2.2464 | -34.046 | -4.0174
2.2525 | -34.0463 | -4.0288
2.2586 | -34.0466 | -4.0393
2.2647 | -34.0469 | -4.0496
2.2708 | -34.0472 | -4.0607
2.2769 | -34.0475 | -4.0713
2.283 | -34.0477 | -4.0823
2.2891 | -34.0479 | -4.0932
2.2952 | -34.048 | -4.1036
2.3013 | -34.0483 | -4.1144
2.3074 | -34.0485 | -4.1247
2.3135 | -34.0487 | -4.136
2.3196 | -34.0489 | -4.146
2.3257 | -34.0491 | -4.1566
2.3318 | -34.0492 | -4.1676
2.3379 | -34.0494 | -4.1782
2.344 | -34.0495 | -4.1896
2.3501 | -34.0497 | -4.2001
2.3562 | -34.0498 | -4.2104
2.3623 | -34.0499 | -4.2206
2.3685 | -34.05 | -4.231
2.3746 | -34.0501 | -4.2422
2.3807 | -34.0503 | -4.2525
2.3868 | -34.0505 | -4.2631
2.3929 | -34.0506 | -4.2736
2.399 | -34.0507 | -4.284
2.4051 | -34.051 | -4.2943
2.4112 | -34.0511 | -4.3048
2.4173 | -34.0513 | -4.3153
2.4234 | -34.0513 | -4.326
2.4295 | -34.0514 | -4.3367
2.4356 | -34.0516 | -4.3475
2.4417 | -34.0517 | -4.3585
2.4478 | -34.0518 | -4.3693
2.4539 | -34.0519 | -4.3793
2.46 | -34.052 | -4.3907
2.4661 | -34.0521 | -4.4017
2.4722 | -34.0522 | -4.4117
2.4783 | -34.0523 | -4.4219
2.4844 | -34.0523 | -4.4329
2.4905 | -34.0524 | -4.4437
2.4966 | -34.0525 | -4.4544
2.5027 | -34.0526 | -4.4647
2.5089 | -34.0527 | -4.4755
2.515 | -34.0527 | -4.4862
2.5211 | -34.0528 | -4.4968
2.5272 | -34.0528 | -4.5076
2.5333 | -34.0529 | -4.5184
2.5394 | -34.053 | -4.5291
2.5455 | -34.0531 | -4.5404
2.5516 | -34.0532 | -4.5506
2.5577 | -34.0534 | -4.5619
2.5638 | -34.0535 | -4.5736
2.5699 | -34.0535 | -4.5849
2.576 | -34.0536 | -4.5964
2.5821 | -34.0536 | -4.6076
2.5882 | -34.0537 | -4.6186
2.5943 | -34.0538 | -4.6293
2.6004 | -34.054 | -4.6402
2.6065 | -34.0541 | -4.6505
2.6126 | -34.0541 | -4.6608
2.6187 | -34.0542 | -4.6711
2.6248 | -34.0543 | -4.6817
2.6309 | -34.0545 | -4.6931
2.637 | -34.0547 | -4.7047
2.6431 | -34.0548 | -4.715
2.6492 | -34.0549 | -4.7263
2.6554 | -34.0551 | -4.737
2.6615 | -34.0551 | -4.7479
2.6676 | -34.0554 | -4.7592
2.6737 | -34.0556 | -4.7703
2.6798 | -34.0559 | -4.7813
2.6859 | -34.056 | -4.7925
2.692 | -34.0562 | -4.8033
2.6981 | -34.0563 | -4.8143
2.7042 | -34.0565 | -4.8254
2.7103 | -34.0567 | -4.8359
2.7164 | -34.057 | -4.8468
2.7225 | -34.0572 | -4.8581
2.7286 | -34.0574 | -4.869
2.7347 | -34.0576 | -4.8794
2.7408 | -34.0579 | -4.8894
2.7469 | -34.0581 | -4.9006
2.753 | -34.0583 | -4.9111
2.7591 | -34.0586 | -4.9223
2.7652 | -34.0588 | -4.9337
2.7713 | -34.0589 | -4.9443
2.7774 | -34.0592 | -4.9552
2.7835 | -34.0594 | -4.9661
2.7896 | -34.0595 | -4.9765
2.7958 | -34.0598 | -4.9871
2.8019 | -34.0601 | -4.9974
2.808 | -34.0603 | -5.008
2.8141 | -34.0606 | -5.0191
2.8202 | -34.0609 | -5.0299
2.8263 | -34.0612 | -5.0407
2.8324 | -34.0616 | -5.052
2.8385 | -34.062 | -5.0627
2.8446 | -34.0623 | -5.0733
2.8507 | -34.0627 | -5.0841
2.8568 | -34.063 | -5.0944
2.8629 | -34.0632 | -5.1049
2.869 | -34.0636 | -5.1159
2.8751 | -34.0639 | -5.1266
2.8812 | -34.0643 | -5.1378
2.8873 | -34.0646 | -5.1486
2.8934 | -34.0649 | -5.1595
2.8995 | -34.0652 | -5.1705
2.9056 | -34.0655 | -5.1816
2.9117 | -34.0658 | -5.1926
2.9178 | -34.0662 | -5.2032
2.9239 | -34.0666 | -5.2147
2.93 | -34.067 | -5.2261
2.9361 | -34.0672 | -5.2368
2.9423 | -34.0676 | -5.2473
2.9484 | -34.068 | -5.2583
2.9545 | -34.0684 | -5.2692
2.9606 | -34.0688 | -5.2805
2.9667 | -34.0691 | -5.2916
2.9728 | -34.0695 | -5.3028
2.9789 | -34.0699 | -5.3138
2.985 | -34.0703 | -5.3251
2.9911 | -34.0707 | -5.3363
2.9972 | -34.0712 | -5.3473
3.0033 | -34.0716 | -5.3588
3.0094 | -34.0719 | -5.3698
3.0155 | -34.0724 | -5.3806
3.0216 | -34.0727 | -5.3918
3.0277 | -34.0731 | -5.4034
3.0338 | -34.0735 | -5.4148
3.0399 | -34.0739 | -5.426
3.046 | -34.0742 | -5.4375
3.0521 | -34.0745 | -5.4486
3.0582 | -34.0748 | -5.4607
3.0643 | -34.0751 | -5.4721
3.0704 | -34.0755 | -5.4839
3.0765 | -34.0757 | -5.4949
3.0827 | -34.076 | -5.5061
3.0888 | -34.0764 | -5.5173
3.0949 | -34.0768 | -5.5286
3.101 | -34.0771 | -5.54
3.1071 | -34.0774 | -5.5518
3.1132 | -34.0778 | -5.5635
3.1193 | -34.0781 | -5.5749
3.1254 | -34.0785 | -5.5862
3.1315 | -34.0789 | -5.5981
3.1376 | -34.0791 | -5.6095
3.1437 | -34.0795 | -5.621
3.1498 | -34.0798 | -5.6331
3.1559 | -34.0801 | -5.6445
3.162 | -34.0804 | -5.6566
3.1681 | -34.0806 | -5.6686
3.1742 | -34.081 | -5.6801
3.1803 | -34.0813 | -5.6917
3.1864 | -34.0816 | -5.703
3.1925 | -34.0819 | -5.7145
3.1986 | -34.0822 | -5.7265
3.2047 | -34.0824 | -5.7386
3.2108 | -34.0827 | -5.7496
3.2169 | -34.083 | -5.7613
3.223 | -34.0832 | -5.7731
3.2292 | -34.0834 | -5.7849
3.2353 | -34.0836 | -5.7957
3.2414 | -34.0839 | -5.8071
3.2475 | -34.084 | -5.8184
3.2536 | -34.0844 | -5.8299
3.2597 | -34.0846 | -5.8408
3.2658 | -34.0848 | -5.8516
3.2719 | -34.085 | -5.8626
3.278 | -34.0852 | -5.8742
3.2841 | -34.0855 | -5.8853
3.2902 | -34.0857 | -5.8967
3.2963 | -34.0859 | -5.9082
3.3024 | -34.086 | -5.9192
3.3085 | -34.0863 | -5.9307
3.3146 | -34.0864 | -5.9418
3.3207 | -34.0867 | -5.9533
3.3268 | -34.0869 | -5.9648
3.3329 | -34.087 | -5.9755
3.339 | -34.0872 | -5.9862
3.3451 | -34.0874 | -5.9974
3.3512 | -34.0875 | -6.008
3.3573 | -34.0877 | -6.0192
3.3634 | -34.0879 | -6.0303
3.3696 | -34.0881 | -6.0416
3.3757 | -34.0882 | -6.0524
3.3818 | -34.0884 | -6.0632
3.3879 | -34.0885 | -6.0742
3.394 | -34.0887 | -6.0856
3.4001 | -34.0888 | -6.0963
3.4062 | -34.0889 | -6.1072
3.4123 | -34.089 | -6.1177
3.4184 | -34.0892 | -6.129
3.4245 | -34.0894 | -6.1397
3.4306 | -34.0895 | -6.1503
3.4367 | -34.0896 | -6.1607
3.4428 | -34.0898 | -6.171
3.4489 | -34.0899 | -6.1819
3.455 | -34.09 | -6.1927
3.4611 | -34.0902 | -6.2039
3.4672 | -34.0903 | -6.2149
3.4733 | -34.0905 | -6.2259
3.4794 | -34.0906 | -6.2367
3.4855 | -34.0907 | -6.247
3.4916 | -34.0909 | -6.2573
3.4977 | -34.091 | -6.2685
3.5038 | -34.0911 | -6.2793
3.5099 | -34.0912 | -6.2895
3.5161 | -34.0914 | -6.2998
3.5222 | -34.0915 | -6.31
3.5283 | -34.0915 | -6.3207
3.5344 | -34.0916 | -6.3314
3.5405 | -34.0918 | -6.3419
3.5466 | -34.092 | -6.3528
3.5527 | -34.092 | -6.3635
3.5588 | -34.0921 | -6.374
3.5649 | -34.0921 | -6.3846
3.571 | -34.0922 | -6.3951
3.5771 | -34.0923 | -6.4062
3.5832 | -34.0925 | -6.4159
3.5893 | -34.0927 | -6.4265
3.5954 | -34.0928 | -6.4372
3.6015 | -34.093 | -6.448
3.6076 | -34.0932 | -6.4585
3.6137 | -34.0933 | -6.4685
3.6198 | -34.0935 | -6.4786
3.6259 | -34.0936 | -6.4894
3.632 | -34.0938 | -6.4999
3.6381 | -34.0939 | -6.5106
3.6442 | -34.0941 | -6.5211
3.6503 | -34.0943 | -6.5312
3.6565 | -34.0945 | -6.5418
3.6626 | -34.0945 | -6.5524
3.6687 | -34.0948 | -6.563
3.6748 | -34.095 | -6.5733
3.6809 | -34.0952 | -6.5839
3.687 | -34.0954 | -6.5948
3.6931 | -34.0956 | -6.6049
3.6992 | -34.0958 | -6.6154
3.7053 | -34.096 | -6.6258
3.7114 | -34.0961 | -6.6364
3.7175 | -34.0963 | -6.647
3.7236 | -34.0966 | -6.6569
3.7297 | -34.0968 | -6.6676
3.7358 | -34.097 | -6.6785
3.7419 | -34.0971 | -6.6893
3.748 | -34.0973 | -6.6997
3.7541 | -34.0975 | -6.7102
3.7602 | -34.0978 | -6.7206
3.7663 | -34.098 | -6.731
3.7724 | -34.0982 | -6.7416
3.7785 | -34.0985 | -6.7519
3.7846 | -34.0987 | -6.7622
3.7907 | -34.099 | -6.7731
3.7969 | -34.0992 | -6.7837
3.803 | -34.0995 | -6.7938
3.8091 | -34.0997 | -6.8042
3.8152 | -34.0999 | -6.8146
3.8213 | -34.1001 | -6.8254
3.8274 | -34.1004 | -6.8356
3.8335 | -34.1007 | -6.8463
3.8396 | -34.101 | -6.8568
3.8457 | -34.1013 | -6.8675
3.8518 | -34.1017 | -6.8779
3.8579 | -34.102 | -6.8884
3.864 | -34.1022 | -6.8991
3.8701 | -34.1026 | -6.9101
3.8762 | -34.1028 | -6.9204
3.8823 | -34.103 | -6.9311
3.8884 | -34.1033 | -6.9417
3.8945 | -34.1037 | -6.9525
3.9006 | -34.104 | -6.9635
3.9067 | -34.1043 | -6.9745
3.9128 | -34.1045 | -6.9852
3.9189 | -34.1048 | -6.9955
3.925 | -34.1051 | -7.0063
3.9311 | -34.1053 | -7.0164
3.9372 | -34.1056 | -7.0272
3.9434 | -34.1059 | -7.0378
3.9495 | -34.1063 | -7.0486
3.9556 | -34.1066 | -7.059
3.9617 | -34.1069 | -7.0692
3.9678 | -34.1072 | -7.0795
3.9739 | -34.1075 | -7.0901
3.98 | -34.1078 | -7.1008
3.9861 | -34.108 | -7.1113
3.9922 | -34.1083 | -7.1219
3.9983 | -34.1086 | -7.1321
4.0044 | -34.1089 | -7.1422
4.0105 | -34.1092 | -7.1528
4.0166 | -34.1095 | -7.1635
4.0227 | -34.1098 | -7.174
4.0288 | -34.11 | -7.1851
4.0349 | -34.1102 | -7.1954
4.041 | -34.1105 | -7.206
4.0471 | -34.1107 | -7.2167
4.0532 | -34.111 | -7.2268
4.0593 | -34.1113 | -7.2376
4.0654 | -34.1116 | -7.2479
4.0715 | -34.1119 | -7.259
4.0776 | -34.1122 | -7.2695
4.0838 | -34.1125 | -7.2805
4.0899 | -34.1128 | -7.2908
4.096 | -34.1131 | -7.3013
4.1021 | -34.1133 | -7.3116
4.1082 | -34.1136 | -7.3219
4.1143 | -34.1138 | -7.3321
4.1204 | -34.1141 | -7.3424
4.1265 | -34.1144 | -7.3525
4.1326 | -34.1147 | -7.3627
4.1387 | -34.115 | -7.3731
4.1448 | -34.1153 | -7.3836
4.1509 | -34.1155 | -7.3942
4.157 | -34.1157 | -7.4051
4.1631 | -34.116 | -7.4158
4.1692 | -34.1163 | -7.4261
4.1753 | -34.1165 | -7.4365
4.1814 | -34.1167 | -7.4466
4.1875 | -34.117 | -7.4574
4.1936 | -34.1173 | -7.468
4.1997 | -34.1175 | -7.4781
4.2058 | -34.1178 | -7.4889
4.2119 | -34.118 | -7.4991
4.218 | -34.1182 | -7.5093
4.2241 | -34.1184 | -7.5195
4.2303 | -34.1187 | -7.5298
4.2364 | -34.119 | -7.5402
4.2425 | -34.1191 | -7.5506
4.2486 | -34.1193 | -7.5612
4.2547 | -34.1195 | -7.5713
4.2608 | -34.1197 | -7.5819
4.2669 | -34.1199 | -7.5919
4.273 | -34.1201 | -7.6024
4.2791 | -34.1203 | -7.6128
4.2852 | -34.1206 | -7.6234
4.2913 | -34.1207 | -7.6334
4.2974 | -34.1209 | -7.6441
4.3035 | -34.1212 | -7.6547
4.3096 | -34.1213 | -7.6653
4.3157 | -34.1216 | -7.6754
4.3218 | -34.1218 | -7.6857
4.3279 | -34.122 | -7.6961
4.334 | -34.1222 | -7.7063
4.3401 | -34.1224 | -7.7162
4.3462 | -34.1226 | -7.7267
4.3523 | -34.1229 | -7.7365
4.3584 | -34.123 | -7.7464
4.3645 | -34.1232 | -7.7566
4.3707 | -34.1234 | -7.7666
4.3768 | -34.1236 | -7.7768
4.3829 | -34.1239 | -7.7875
4.389 | -34.124 | -7.7978
4.3951 | -34.1242 | -7.8078
4.4012 | -34.1244 | -7.8184
4.4073 | -34.1246 | -7.8285
4.4134 | -34.1248 | -7.8387
4.4195 | -34.125 | -7.8489
4.4256 | -34.1253 | -7.8589
4.4317 | -34.1255 | -7.869
4.4378 | -34.1257 | -7.8789
4.4439 | -34.1258 | -7.8888
4.45 | -34.126 | -7.899
4.4561 | -34.1261 | -7.9091
4.4622 | -34.1263 | -7.9194
4.4683 | -34.1265 | -7.9295
4.4744 | -34.1267 | -7.9401
4.4805 | -34.1269 | -7.9498
4.4866 | -34.1271 | -7.96
4.4927 | -34.1273 | -7.9704
4.4988 | -34.1275 | -7.9805
4.5049 | -34.1277 | -7.9903
4.511 | -34.1279 | -8.0001
4.5172 | -34.1281 | -8.0102
4.5233 | -34.1282 | -8.0208
4.5294 | -34.1285 | -8.0308
4.5355 | -34.1287 | -8.0409
4.5416 | -34.1288 | -8.051
4.5477 | -34.129 | -8.0611
4.5538 | -34.1292 | -8.0708
4.5599 | -34.1294 | -8.0809
4.566 | -34.1296 | -8.091
4.5721 | -34.1298 | -8.1012
4.5782 | -34.1301 | -8.1116
4.5843 | -34.1303 | -8.1218
4.5904 | -34.1305 | -8.1319
4.5965 | -34.1307 | -8.1419
4.6026 | -34.1309 | -8.1525
4.6087 | -34.1311 | -8.1625
4.6148 | -34.1313 | -8.173
4.6209 | -34.1315 | -8.1831
4.627 | -34.1317 | -8.1934
4.6331 | -34.1319 | -8.2034
4.6392 | -34.1321 | -8.2137
4.6453 | -34.1323 | -8.2238
4.6514 | -34.1325 | -8.234
4.6576 | -34.1327 | -8.2445
4.6637 | -34.1328 | -8.2542
4.6698 | -34.133 | -8.2642
4.6759 | -34.1332 | -8.2747
4.682 | -34.1334 | -8.2847
4.6881 | -34.1335 | -8.2948
4.6942 | -34.1337 | -8.3049
4.7003 | -34.1339 | -8.3151
4.7064 | -34.1341 | -8.3251
4.7125 | -34.1344 | -8.3355
4.7186 | -34.1345 | -8.3454
4.7247 | -34.1348 | -8.3556
4.7308 | -34.135 | -8.3657
4.7369 | -34.1352 | -8.3759
4.743 | -34.1354 | -8.3865
4.7491 | -34.1356 | -8.397
4.7552 | -34.1358 | -8.4072
4.7613 | -34.1359 | -8.4175
4.7674 | -34.1361 | -8.4276
4.7735 | -34.1363 | -8.4382
4.7796 | -34.1366 | -8.4487
4.7857 | -34.1368 | -8.459
4.7918 | -34.137 | -8.469
4.7979 | -34.1372 | -8.479
4.8041 | -34.1374 | -8.4893
4.8102 | -34.1376 | -8.4991
4.8163 | -34.1378 | -8.5092
4.8224 | -34.138 | -8.5195
4.8285 | -34.1383 | -8.5293
4.8346 | -34.1385 | -8.5393
4.8407 | -34.1387 | -8.5496
4.8468 | -34.1389 | -8.5601
4.8529 | -34.1392 | -8.5706
4.859 | -34.1394 | -8.5811
4.8651 | -34.1396 | -8.591
4.8712 | -34.1398 | -8.601
4.8773 | -34.14 | -8.6113
4.8834 | -34.1402 | -8.6216
4.8895 | -34.1404 | -8.6315
4.8956 | -34.1406 | -8.6417
4.9017 | -34.1408 | -8.6517
4.9078 | -34.141 | -8.6623
4.9139 | -34.1412 | -8.672
4.92 | -34.1415 | -8.6819
4.9261 | -34.1417 | -8.6917
4.9322 | -34.142 | -8.702
4.9383 | -34.1422 | -8.7122
4.9445 | -34.1424 | -8.7223
4.9506 | -34.1426 | -8.7323
4.9567 | -34.1429 | -8.7423
4.9628 | -34.1431 | -8.7524
4.9689 | -34.1433 | -8.763
4.975 | -34.1435 | -8.7735
4.9811 | -34.1437 | -8.7834
4.9872 | -34.144 | -8.7934
4.9933 | -34.1443 | -8.804
4.9994 | -34.1446 | -8.8146
5.0055 | -34.1448 | -8.8248
5.0116 | -34.1451 | -8.8355
5.0177 | -34.1453 | -8.846
5.0238 | -34.1456 | -8.8565
5.0299 | -34.1459 | -8.8671
5.036 | -34.1462 | -8.8774
5.0421 | -34.1464 | -8.8875
5.0482 | -34.1466 | -8.8978
5.0543 | -34.1469 | -8.9078
5.0604 | -34.1472 | -8.918
5.0665 | -34.1474 | -8.9279
5.0726 | -34.1477 | -8.9385
5.0787 | -34.1479 | -8.949
5.0848 | -34.1482 | -8.9593
5.091 | -34.1484 | -8.9693
5.0971 | -34.1487 | -8.9793
5.1032 | -34.149 | -8.9899
5.1093 | -34.1493 | -9.0001
5.1154 | -34.1495 | -9.0101
5.1215 | -34.1497 | -9.0201
5.1276 | -34.15 | -9.0302
5.1337 | -34.1503 | -9.0401
5.1398 | -34.1505 | -9.0501
5.1459 | -34.1508 | -9.0606
5.152 | -34.151 | -9.071
5.1581 | -34.1513 | -9.0815
5.1642 | -34.1516 | -9.0918
5.1703 | -34.1518 | -9.1021
5.1764 | -34.1521 | -9.1122
5.1825 | -34.1524 | -9.1221
5.1886 | -34.1527 | -9.1321
5.1947 | -34.1529 | -9.1424
5.2008 | -34.1532 | -9.1527
5.2069 | -34.1534 | -9.1632
5.213 | -34.1537 | -9.1733
5.2191 | -34.154 | -9.1836
5.2252 | -34.1543 | -9.1939
5.2314 | -34.1545 | -9.2043
5.2375 | -34.1548 | -9.2143
5.2436 | -34.155 | -9.2247
5.2497 | -34.1553 | -9.2346
5.2558 | -34.1556 | -9.2453
5.2619 | -34.1559 | -9.2558
5.268 | -34.1562 | -9.2663
5.2741 | -34.1565 | -9.2766
5.2802 | -34.1567 | -9.2873
5.2863 | -34.1571 | -9.2975
5.2924 | -34.1573 | -9.308
5.2985 | -34.1577 | -9.3184
5.3046 | -34.1579 | -9.3288
5.3107 | -34.1582 | -9.3393
5.3168 | -34.1585 | -9.3497
5.3229 | -34.1588 | -9.36
5.329 | -34.159 | -9.3704
5.3351 | -34.1593 | -9.3803
5.3412 | -34.1596 | -9.3906
5.3473 | -34.1599 | -9.401
5.3534 | -34.1602 | -9.4115
5.3595 | -34.1605 | -9.4219
5.3656 | -34.1608 | -9.4322
5.3717 | -34.1611 | -9.4429
5.3779 | -34.1613 | -9.4536
5.384 | -34.1616 | -9.4639
5.3901 | -34.1619 | -9.4746
5.3962 | -34.1622 | -9.4847
5.4023 | -34.1625 | -9.4953
5.4084 | -34.1628 | -9.5056
5.4145 | -34.1631 | -9.5157
5.4206 | -34.1634 | -9.5261
5.4267 | -34.1637 | -9.5368
5.4328 | -34.164 | -9.5473
5.4389 | -34.1643 | -9.5574
5.445 | -34.1645 | -9.5681
5.4511 | -34.1649 | -9.5783
5.4572 | -34.1652 | -9.5887
5.4633 | -34.1654 | -9.5991
5.4694 | -34.1657 | -9.6095
5.4755 | -34.166 | -9.6196
5.4816 | -34.1663 | -9.63
5.4877 | -34.1666 | -9.6403
5.4938 | -34.1669 | -9.6503
5.4999 | -34.1672 | -9.6606
5.506 | -34.1675 | -9.671
5.5121 | -34.1678 | -9.6812
5.5183 | -34.1681 | -9.6914
5.5244 | -34.1683 | -9.702
5.5305 | -34.1687 | -9.7123
5.5366 | -34.169 | -9.7224
5.5427 | -34.1692 | -9.7329
5.5488 | -34.1695 | -9.7434
5.5549 | -34.1698 | -9.7538
5.561 | -34.1701 | -9.7644
5.5671 | -34.1704 | -9.7746
5.5732 | -34.1707 | -9.7853
5.5793 | -34.171 | -9.796
5.5854 | -34.1713 | -9.8063
5.5915 | -34.1716 | -9.8171
5.5976 | -34.1719 | -9.8275
5.6037 | -34.1722 | -9.838
5.6098 | -34.1724 | -9.8483
5.6159 | -34.1727 | -9.8585
5.622 | -34.173 | -9.869
5.6281 | -34.1733 | -9.8795
5.6342 | -34.1736 | -9.8903
5.6403 | -34.1739 | -9.9007
5.6464 | -34.1742 | -9.9113
5.6525 | -34.1745 | -9.9218
5.6586 | -34.1748 | -9.9319
5.6648 | -34.1751 | -9.9423
5.6709 | -34.1755 | -9.9529
5.677 | -34.1758 | -9.9634
5.6831 | -34.1761 | -9.9738
5.6892 | -34.1763 | -9.9845
5.6953 | -34.1766 | -9.9949
5.7014 | -34.1769 | -10.0055
5.7075 | -34.1772 | -10.0158
5.7136 | -34.1775 | -10.0262
5.7197 | -34.1778 | -10.0366
5.7258 | -34.178 | -10.0468
5.7319 | -34.1784 | -10.0573
5.738 | -34.1787 | -10.0677
5.7441 | -34.179 | -10.0781
5.7502 | -34.1793 | -10.0882
5.7563 | -34.1796 | -10.0986
5.7624 | -34.1799 | -10.1093
5.7685 | -34.1802 | -10.1197
5.7746 | -34.1805 | -10.1305
5.7807 | -34.1808 | -10.1409
5.7868 | -34.181 | -10.1516
5.7929 | -34.1813 | -10.1619
5.799 | -34.1816 | -10.1722
5.8052 | -34.1819 | -10.1824
5.8113 | -34.1822 | -10.1928
5.8174 | -34.1825 | -10.203
5.8235 | -34.1828 | -10.2134
5.8296 | -34.183 | -10.224
5.8357 | -34.1833 | -10.2344
5.8418 | -34.1836 | -10.245
5.8479 | -34.1839 | -10.2555
5.854 | -34.1842 | -10.266
5.8601 | -34.1845 | -10.2763
5.8662 | -34.1848 | -10.2869
5.8723 | -34.185 | -10.2971
5.8784 | -34.1853 | -10.3075
5.8845 | -34.1856 | -10.3177
5.8906 | -34.1859 | -10.3282
5.8967 | -34.1861 | -10.3384
5.9028 | -34.1865 | -10.3487
5.9089 | -34.1868 | -10.3589
5.915 | -34.1871 | -10.3691
5.9211 | -34.1874 | -10.3794
5.9272 | -34.1877 | -10.3895
5.9333 | -34.188 | -10.3998
5.9394 | -34.1883 | -10.4102
5.9455 | -34.1886 | -10.4204
5.9517 | -34.1889 | -10.4305
5.9578 | -34.1893 | -10.441
5.9639 | -34.1895 | -10.451
5.97 | -34.1899 | -10.4615
5.9761 | -34.1902 | -10.4718
5.9822 | -34.1905 | -10.4822
5.9883 | -34.1908 | -10.4925
5.9944 | -34.1911 | -10.503
6.0005 | -34.1914 | -10.5135
6.0066 | -34.1917 | -10.5237
6.0127 | -34.192 | -10.5337
6.0188 | -34.1923 | -10.5444
6.0249 | -34.1926 | -10.5548
6.031 | -34.1929 | -10.565
6.0371 | -34.1932 | -10.5753
6.0432 | -34.1935 | -10.5855
6.0493 | -34.1938 | -10.5956
6.0554 | -34.1941 | -10.6058
6.0615 | -34.1944 | -10.6161
6.0676 | -34.1948 | -10.6267
6.0737 | -34.1951 | -10.6371
6.0798 | -34.1953 | -10.6474
6.0859 | -34.1957 | -10.6577
6.0921 | -34.196 | -10.6677
6.0982 | -34.1963 | -10.678
6.1043 | -34.1966 | -10.6882
6.1104 | -34.197 | -10.6987
6.1165 | -34.1973 | -10.7085
6.1226 | -34.1976 | -10.7187
6.1287 | -34.1979 | -10.729
6.1348 | -34.1982 | -10.7392
6.1409 | -34.1985 | -10.7495
6.147 | -34.1988 | -10.7597
6.1531 | -34.1991 | -10.7699
6.1592 | -34.1994 | -10.78
6.1653 | -34.1997 | -10.7901
6.1714 | -34.2 | -10.8004
6.1775 | -34.2003 | -10.8108
6.1836 | -34.2006 | -10.8211
6.1897 | -34.2009 | -10.8313
6.1958 | -34.2012 | -10.8418
6.2019 | -34.2016 | -10.8522
6.208 | -34.2019 | -10.8624
6.2141 | -34.2022 | -10.8726
6.2202 | -34.2025 | -10.883
6.2263 | -34.2028 | -10.8931
6.2325 | -34.2031 | -10.903
6.2386 | -34.2034 | -10.9129
6.2447 | -34.2038 | -10.9234
6.2508 | -34.2041 | -10.9336
6.2569 | -34.2044 | -10.9438
6.263 | -34.2047 | -10.954
6.2691 | -34.205 | -10.9641
6.2752 | -34.2053 | -10.9744
6.2813 | -34.2056 | -10.9846
6.2874 | -34.2059 | -10.9946
6.2935 | -34.2062 | -11.0048
6.2996 | -34.2065 | -11.0149
6.3057 | -34.2068 | -11.0252
6.3118 | -34.207 | -11.0355
6.3179 | -34.2073 | -11.0457
6.324 | -34.2077 | -11.0558
6.3301 | -34.208 | -11.0659
6.3362 | -34.2083 | -11.076
6.3423 | -34.2086 | -11.0864
6.3484 | -34.2089 | -11.0966
6.3545 | -34.2092 | -11.1065
6.3606 | -34.2095 | -11.1164
6.3667 | -34.2098 | -11.1264
6.3728 | -34.2101 | -11.1366
6.379 | -34.2104 | -11.1468
6.3851 | -34.2107 | -11.1567
6.3912 | -34.211 | -11.1667
6.3973 | -34.2113 | -11.1769
6.4034 | -34.2116 | -11.1868
6.4095 | -34.2119 | -11.197
6.4156 | -34.2122 | -11.2071
6.4217 | -34.2125 | -11.2172
6.4278 | -34.2128 | -11.2275
6.4339 | -34.2131 | -11.2375
6.44 | -34.2134 | -11.2475
6.4461 | -34.2137 | -11.2579
6.4522 | -34.214 | -11.2677
6.4583 | -34.2143 | -11.278
6.4644 | -34.2146 | -11.2882
6.4705 | -34.2149 | -11.298
6.4766 | -34.2152 | -11.3081
6.4827 | -34.2155 | -11.3182
6.4888 | -34.2157 | -11.3281
6.4949 | -34.216 | -11.3379
6.501 | -34.2163 | -11.3481
6.5071 | -34.2167 | -11.3579
6.5132 | -34.217 | -11.3681
6.5194 | -34.2172 | -11.3782
6.5255 | -34.2176 | -11.388
6.5316 | -34.2178 | -11.3983
6.5377 | -34.2181 | -11.4083
6.5438 | -34.2184 | -11.4181
6.5499 | -34.2187 | -11.4281
6.556 | -34.219 | -11.4381
6.5621 | -34.2193 | -11.4484
6.5682 | -34.2196 | -11.4585
6.5743 | -34.2199 | -11.4684
6.5804 | -34.2202 | -11.4786
6.5865 | -34.2205 | -11.4886
6.5926 | -34.2208 | -11.4986
6.5987 | -34.2211 | -11.5083
6.6048 | -34.2214 | -11.5184
6.6109 | -34.2217 | -11.5284
6.617 | -34.222 | -11.5386
6.6231 | -34.2223 | -11.5485
6.6292 | -34.2226 | -11.5586
6.6353 | -34.2229 | -11.5689
6.6414 | -34.2232 | -11.5788
6.6475 | -34.2235 | -11.5888
6.6536 | -34.2238 | -11.5987
6.6597 | -34.2241 | -11.6084
6.6659 | -34.2244 | -11.6185
6.672 | -34.2247 | -11.6283
6.6781 | -34.225 | -11.6383
6.6842 | -34.2253 | -11.6482
6.6903 | -34.2257 | -11.6581
6.6964 | -34.226 | -11.6682
6.7025 | -34.2263 | -11.6779
6.7086 | -34.2266 | -11.6879
6.7147 | -34.2269 | -11.698
6.7208 | -34.2271 | -11.7081
6.7269 | -34.2274 | -11.718
6.733 | -34.2277 | -11.7277
6.7391 | -34.228 | -11.7377
6.7452 | -34.2284 | -11.7474
6.7513 | -34.2287 | -11.757
6.7574 | -34.229 | -11.7671
6.7635 | -34.2293 | -11.7768
6.7696 | -34.2296 | -11.7866
6.7757 | -34.2299 | -11.7966
6.7818 | -34.2302 | -11.8066
6.7879 | -34.2306 | -11.8163
6.794 | -34.2308 | -11.826
6.8001 | -34.2311 | -11.8359
6.8063 | -34.2314 | -11.8458
6.8124 | -34.2317 | -11.8559
6.8185 | -34.232 | -11.866
6.8246 | -34.2323 | -11.8756
6.8307 | -34.2326 | -11.8854
6.8368 | -34.2329 | -11.8956
6.8429 | -34.2332 | -11.9055
6.849 | -34.2335 | -11.9155
6.8551 | -34.2339 | -11.9252
6.8612 | -34.2341 | -11.9351
6.8673 | -34.2344 | -11.9449
6.8734 | -34.2347 | -11.9549
6.8795 | -34.235 | -11.9648
6.8856 | -34.2353 | -11.9747
6.8917 | -34.2356 | -11.9846
6.8978 | -34.2359 | -11.9942
6.9039 | -34.2362 | -12.0038
6.91 | -34.2366 | -12.0137
6.9161 | -34.2369 | -12.0233
6.9222 | -34.2372 | -12.0332
6.9283 | -34.2375 | -12.0427
6.9344 | -34.2378 | -12.0527
6.9405 | -34.2382 | -12.0625
6.9466 | -34.2385 | -12.0725
6.9528 | -34.2388 | -12.0826
6.9589 | -34.2392 | -12.0926
6.965 | -34.2395 | -12.1025
6.9711 | -34.2398 | -12.1124
6.9772 | -34.2402 | -12.1223
6.9833 | -34.2405 | -12.1321
6.9894 | -34.2408 | -12.1422
6.9955 | -34.2411 | -12.1518
7.0016 | -34.2414 | -12.1616
7.0077 | -34.2417 | -12.1714
7.0138 | -34.2421 | -12.1812
7.0199 | -34.2423 | -12.1909
7.026 | -34.2427 | -12.2007
7.0321 | -34.243 | -12.2103
7.0382 | -34.2433 | -12.2198
7.0443 | -34.2437 | -12.2298
7.0504 | -34.244 | -12.2397
7.0565 | -34.2443 | -12.2495
7.0626 | -34.2446 | -12.2591
7.0687 | -34.2449 | -12.2685
7.0748 | -34.2453 | -12.278
7.0809 | -34.2456 | -12.2878
7.087 | -34.2459 | -12.2976
7.0932 | -34.2462 | -12.3075
7.0993 | -34.2465 | -12.3172
7.1054 | -34.2469 | -12.3268
7.1115 | -34.2472 | -12.3364
7.1176 | -34.2475 | -12.3461
7.1237 | -34.2478 | -12.3558
7.1298 | -34.2481 | -12.3655
7.1359 | -34.2484 | -12.3752
7.142 | -34.2487 | -12.3847
7.1481 | -34.2491 | -12.3944
7.1542 | -34.2494 | -12.4041
7.1603 | -34.2497 | -12.4139
7.1664 | -34.2501 | -12.4236
7.1725 | -34.2504 | -12.4334
7.1786 | -34.2507 | -12.4431
7.1847 | -34.251 | -12.4529
7.1908 | -34.2513 | -12.4628
7.1969 | -34.2516 | -12.4728
7.203 | -34.252 | -12.4823
7.2091 | -34.2523 | -12.4919
7.2152 | -34.2526 | -12.5017
7.2213 | -34.253 | -12.5112
7.2274 | -34.2533 | -12.5209
7.2335 | -34.2536 | -12.5306
7.2397 | -34.254 | -12.5403
7.2458 | -34.2543 | -12.5497
7.2519 | -34.2546 | -12.5595
7.258 | -34.255 | -12.5694
7.2641 | -34.2553 | -12.5791
7.2702 | -34.2556 | -12.589
7.2763 | -34.2559 | -12.5987
7.2824 | -34.2563 | -12.6086
7.2885 | -34.2566 | -12.6184
7.2946 | -34.2569 | -12.6281
7.3007 | -34.2572 | -12.6379
7.3068 | -34.2575 | -12.6479
7.3129 | -34.2579 | -12.6578
7.319 | -34.2582 | -12.6676
7.3251 | -34.2585 | -12.6771
7.3312 | -34.2588 | -12.6868
7.3373 | -34.2592 | -12.6963
7.3434 | -34.2596 | -12.7059
7.3495 | -34.2599 | -12.7155
7.3556 | -34.2602 | -12.7253
7.3617 | -34.2606 | -12.7353
7.3678 | -34.2609 | -12.7449
7.3739 | -34.2612 | -12.7549
7.3801 | -34.2615 | -12.7648
7.3862 | -34.2619 | -12.7745
7.3923 | -34.2622 | -12.7844
7.3984 | -34.2625 | -12.7939
7.4045 | -34.2629 | -12.8036
7.4106 | -34.2633 | -12.8132
7.4167 | -34.2636 | -12.823
7.4228 | -34.2639 | -12.8329
7.4289 | -34.2643 | -12.8426
7.435 | -34.2646 | -12.8521
7.4411 | -34.2649 | -12.8619
7.4472 | -34.2653 | -12.8717
7.4533 | -34.2656 | -12.8813
7.4594 | -34.266 | -12.8913
7.4655 | -34.2663 | -12.9008
7.4716 | -34.2666 | -12.9107
7.4777 | -34.2669 | -12.9204
7.4838 | -34.2673 | -12.93
7.4899 | -34.2676 | -12.9397
7.496 | -34.2679 | -12.9493
7.5021 | -34.2683 | -12.9589
7.5082 | -34.2686 | -12.9686
7.5143 | -34.269 | -12.9784
7.5204 | -34.2693 | -12.9882
7.5266 | -34.2696 | -12.9978
7.5327 | -34.27 | -13.0075
7.5388 | -34.2703 | -13.0175
7.5449 | -34.2706 | -13.0274
7.551 | -34.271 | -13.037
7.5571 | -34.2713 | -13.0468
7.5632 | -34.2717 | -13.0567
7.5693 | -34.2721 | -13.0663
7.5754 | -34.2724 | -13.0759
7.5815 | -34.2727 | -13.0857
7.5876 | -34.273 | -13.0956
7.5937 | -34.2734 | -13.1055
7.5998 | -34.2738 | -13.1152
7.6059 | -34.2742 | -13.1251
7.612 | -34.2745 | -13.1348
7.6181 | -34.2748 | -13.1445
7.6242 | -34.2752 | -13.1545
7.6303 | -34.2755 | -13.1645
7.6364 | -34.2758 | -13.1743
7.6425 | -34.2762 | -13.1841
7.6486 | -34.2765 | -13.1939
7.6547 | -34.2768 | -13.2036
7.6608 | -34.2772 | -13.2133
7.667 | -34.2775 | -13.2232
7.6731 | -34.2779 | -13.233
7.6792 | -34.2782 | -13.243
7.6853 | -34.2786 | -13.2529
7.6914 | -34.2789 | -13.2628
7.6975 | -34.2793 | -13.2726
7.7036 | -34.2796 | -13.2825
7.7097 | -34.28 | -13.2924
7.7158 | -34.2803 | -13.3023
7.7219 | -34.2807 | -13.3119
7.728 | -34.281 | -13.3219
7.7341 | -34.2813 | -13.3318
7.7402 | -34.2817 | -13.3417
7.7463 | -34.2821 | -13.3515
7.7524 | -34.2824 | -13.3616
7.7585 | -34.2827 | -13.3714
7.7646 | -34.2831 | -13.3812
7.7707 | -34.2834 | -13.3913
7.7768 | -34.2837 | -13.4013
7.7829 | -34.2841 | -13.4114
7.789 | -34.2844 | -13.4215
7.7951 | -34.2847 | -13.4313
7.8012 | -34.2851 | -13.4411
7.8073 | -34.2855 | -13.4512
7.8135 | -34.2858 | -13.461
7.8196 | -34.2862 | -13.4711
7.8257 | -34.2865 | -13.4813
7.8318 | -34.2868 | -13.4912
7.8379 | -34.2872 | -13.501
7.844 | -34.2876 | -13.5108
7.8501 | -34.2879 | -13.5209
7.8562 | -34.2883 | -13.531
7.8623 | -34.2886 | -13.5408
7.8684 | -34.289 | -13.5508
7.8745 | -34.2893 | -13.5607
7.8806 | -34.2896 | -13.5709
7.8867 | -34.29 | -13.5805
7.8928 | -34.2904 | -13.5901
7.8989 | -34.2907 | -13.6
7.905 | -34.2911 | -13.61
7.9111 | -34.2914 | -13.62
7.9172 | -34.2918 | -13.6299
7.9233 | -34.2921 | -13.6399
7.9294 | -34.2924 | -13.6498
7.9355 | -34.2928 | -13.6595
7.9416 | -34.2931 | -13.6695
7.9477 | -34.2934 | -13.6792
7.9539 | -34.2938 | -13.6891
7.96 | -34.2941 | -13.6991
7.9661 | -34.2945 | -13.7091
7.9722 | -34.2949 | -13.7188
7.9783 | -34.2952 | -13.7289
7.9844 | -34.2956 | -13.7388
7.9905 | -34.2959 | -13.7489
7.9966 | -34.2963 | -13.759
8.0027 | -34.2966 | -13.7689
8.0088 | -34.297 | -13.7788
8.0149 | -34.2973 | -13.7887
8.021 | -34.2976 | -13.7987
8.0271 | -34.298 | -13.809
8.0332 | -34.2983 | -13.8189
8.0393 | -34.2987 | -13.829
8.0454 | -34.2991 | -13.8392
8.0515 | -34.2995 | -13.8491
8.0576 | -34.2998 | -13.8594
8.0637 | -34.3001 | -13.8692
8.0698 | -34.3005 | -13.879
8.0759 | -34.3008 | -13.8888
8.082 | -34.3012 | -13.8986
8.0881 | -34.3015 | -13.9087
8.0942 | -34.3019 | -13.9187
8.1004 | -34.3023 | -13.9287
8.1065 | -34.3026 | -13.9385
8.1126 | -34.3029 | -13.9484
8.1187 | -34.3033 | -13.9586
8.1248 | -34.3036 | -13.9683
8.1309 | -34.304 | -13.9781
8.137 | -34.3043 | -13.9878
8.1431 | -34.3047 | -13.998
8.1492 | -34.305 | -14.0079
8.1553 | -34.3054 | -14.0179
8.1614 | -34.3057 | -14.028
8.1675 | -34.306 | -14.0379
8.1736 | -34.3064 | -14.048
8.1797 | -34.3067 | -14.058
8.1858 | -34.3071 | -14.0681
8.1919 | -34.3074 | -14.078
8.198 | -34.3078 | -14.0878
8.2041 | -34.3082 | -14.0978
8.2102 | -34.3085 | -14.1078
8.2163 | -34.3089 | -14.1175
8.2224 | -34.3092 | -14.1274
8.2285 | -34.3096 | -14.1371
8.2346 | -34.3099 | -14.147
8.2408 | -34.3103 | -14.1568
8.2469 | -34.3106 | -14.1668
8.253 | -34.311 | -14.1765
8.2591 | -34.3114 | -14.1865
8.2652 | -34.3117 | -14.1965
8.2713 | -34.3121 | -14.2065
8.2774 | -34.3125 | -14.2166
8.2835 | -34.3128 | -14.2268
8.2896 | -34.3132 | -14.2365
8.2957 | -34.3135 | -14.2466
8.3018 | -34.3139 | -14.2565
8.3079 | -34.3143 | -14.2663
8.314 | -34.3146 | -14.2762
8.3201 | -34.315 | -14.286
8.3262 | -34.3153 | -14.296
8.3323 | -34.3157 | -14.3059
8.3384 | -34.316 | -14.316
8.3445 | -34.3164 | -14.3258
8.3506 | -34.3168 | -14.3359
8.3567 | -34.3171 | -14.3459
8.3628 | -34.3175 | -14.3558
8.3689 | -34.3178 | -14.3656
8.375 | -34.3182 | -14.3756
8.3812 | -34.3185 | -14.3854
8.3873 | -34.3189 | -14.3954
8.3934 | -34.3192 | -14.4054
8.3995 | -34.3196 | -14.4152
8.4056 | -34.3199 | -14.425
8.4117 | -34.3203 | -14.435
8.4178 | -34.3206 | -14.445
8.4239 | -34.321 | -14.4547
8.43 | -34.3213 | -14.4647
8.4361 | -34.3216 | -14.4747
8.4422 | -34.322 | -14.4845
8.4483 | -34.3224 | -14.4943
8.4544 | -34.3228 | -14.5043
8.4605 | -34.3231 | -14.5141
8.4666 | -34.3234 | -14.5242
8.4727 | -34.3238 | -14.5342
8.4788 | -34.3241 | -14.544
8.4849 | -34.3245 | -14.5539
8.491 | -34.3249 | -14.5638
8.4971 | -34.3252 | -14.5735
8.5032 | -34.3256 | -14.5832
8.5093 | -34.3259 | -14.5929
8.5154 | -34.3263 | -14.6029
8.5215 | -34.3267 | -14.6127
8.5277 | -34.327 | -14.6226
8.5338 | -34.3273 | -14.6324
8.5399 | -34.3277 | -14.6423
8.546 | -34.3281 | -14.6519
8.5521 | -34.3284 | -14.6618
8.5582 | -34.3288 | -14.6716
8.5643 | -34.3291 | -14.6812
8.5704 | -34.3295 | -14.6908
8.5765 | -34.3298 | -14.7008
8.5826 | -34.3302 | -14.7103
8.5887 | -34.3306 | -14.7201
8.5948 | -34.331 | -14.7299
8.6009 | -34.3313 | -14.7396
8.607 | -34.3317 | -14.7496
8.6131 | -34.332 | -14.7594
8.6192 | -34.3324 | -14.769
8.6253 | -34.3328 | -14.7788
8.6314 | -34.3331 | -14.7885
8.6375 | -34.3335 | -14.7984
8.6436 | -34.3338 | -14.8082
8.6497 | -34.3342 | -14.8179
8.6558 | -34.3345 | -14.8276
8.6619 | -34.3349 | -14.8375
8.6681 | -34.3352 | -14.8474
8.6742 | -34.3356 | -14.8571
8.6803 | -34.336 | -14.8669
8.6864 | -34.3363 | -14.8764
8.6925 | -34.3367 | -14.886
8.6986 | -34.337 | -14.8959
8.7047 | -34.3374 | -14.9058
8.7108 | -34.3377 | -14.9156
8.7169 | -34.3381 | -14.9254
8.723 | -34.3384 | -14.935
8.7291 | -34.3388 | -14.9446
8.7352 | -34.3392 | -14.9546
8.7413 | -34.3395 | -14.9644
8.7474 | -34.3399 | -14.974
8.7535 | -34.3403 | -14.9837
8.7596 | -34.3407 | -14.9935
8.7657 | -34.341 | -15.0033
8.7718 | -34.3413 | -15.0131
8.7779 | -34.3417 | -15.023
8.784 | -34.3421 | -15.0328
8.7901 | -34.3425 | -15.0425
8.7962 | -34.3428 | -15.052
8.8023 | -34.3431 | -15.0616
8.8084 | -34.3435 | -15.0713
8.8146 | -34.3439 | -15.081
8.8207 | -34.3442 | -15.0907
8.8268 | -34.3446 | -15.1
8.8329 | -34.345 | -15.1098
8.839 | -34.3453 | -15.1195
8.8451 | -34.3457 | -15.1291
8.8512 | -34.3461 | -15.1387
8.8573 | -34.3464 | -15.1485
8.8634 | -34.3468 | -15.1579
8.8695 | -34.3472 | -15.1678
8.8756 | -34.3475 | -15.1775
8.8817 | -34.3478 | -15.1871
8.8878 | -34.3482 | -15.197
8.8939 | -34.3486 | -15.2067
8.9 | -34.349 | -15.2164
8.9061 | -34.3493 | -15.2261
8.9122 | -34.3497 | -15.2357
8.9183 | -34.35 | -15.2455
8.9244 | -34.3504 | -15.2553
8.9305 | -34.3508 | -15.2648
8.9366 | -34.3512 | -15.2745
8.9427 | -34.3515 | -15.2842
8.9488 | -34.3519 | -15.2938
8.955 | -34.3523 | -15.3034
8.9611 | -34.3527 | -15.3132
8.9672 | -34.3531 | -15.3228
8.9733 | -34.3534 | -15.3323
8.9794 | -34.3538 | -15.3418
8.9855 | -34.3542 | -15.3516
8.9916 | -34.3545 | -15.3614
8.9977 | -34.3549 | -15.3711
9.0038 | -34.3553 | -15.3805
9.0099 | -34.3557 | -15.3902
9.016 | -34.3561 | -15.3999
9.0221 | -34.3564 | -15.4093
9.0282 | -34.3568 | -15.4188
9.0343 | -34.3572 | -15.4285
9.0404 | -34.3576 | -15.4381
9.0465 | -34.358 | -15.4479
9.0526 | -34.3584 | -15.4573
9.0587 | -34.3588 | -15.467
9.0648 | -34.3592 | -15.4765
9.0709 | -34.3595 | -15.4861
9.077 | -34.3599 | -15.4958
9.0831 | -34.3603 | -15.5056
9.0892 | -34.3607 | -15.5152
9.0953 | -34.3611 | -15.525
9.1015 | -34.3615 | -15.5346
9.1076 | -34.3619 | -15.544
9.1137 | -34.3622 | -15.5536
9.1198 | -34.3626 | -15.5632
9.1259 | -34.363 | -15.5728
9.132 | -34.3634 | -15.5825
9.1381 | -34.3638 | -15.592
9.1442 | -34.3642 | -15.6012
9.1503 | -34.3646 | -15.6109
9.1564 | -34.3649 | -15.6206
9.1625 | -34.3653 | -15.6302
9.1686 | -34.3657 | -15.6397
9.1747 | -34.3661 | -15.6493
9.1808 | -34.3665 | -15.6589
9.1869 | -34.3669 | -15.6684
9.193 | -34.3673 | -15.678
9.1991 | -34.3677 | -15.6877
9.2052 | -34.3681 | -15.6973
9.2113 | -34.3685 | -15.7069
9.2174 | -34.3689 | -15.7163
9.2235 | -34.3693 | -15.7259
9.2296 | -34.3697 | -15.7355
9.2357 | -34.3701 | -15.7452
9.2419 | -34.3705 | -15.7547
9.248 | -34.3709 | -15.7642
9.2541 | -34.3713 | -15.7739
9.2602 | -34.3717 | -15.7835
9.2663 | -34.3721 | -15.7928
9.2724 | -34.3725 | -15.8024
9.2785 | -34.3729 | -15.8119
9.2846 | -34.3733 | -15.8213
9.2907 | -34.3737 | -15.8309
9.2968 | -34.3741 | -15.8407
9.3029 | -34.3745 | -15.8504
9.309 | -34.3748 | -15.8601
9.3151 | -34.3752 | -15.8694
9.3212 | -34.3756 | -15.8787
9.3273 | -34.376 | -15.8882
9.3334 | -34.3764 | -15.8977
9.3395 | -34.3768 | -15.9072
9.3456 | -34.3772 | -15.9166
9.3517 | -34.3775 | -15.9263
9.3578 | -34.378 | -15.9357
9.3639 | -34.3784 | -15.9451
9.37 | -34.3788 | -15.9547
9.3761 | -34.3792 | -15.9644
9.3822 | -34.3796 | -15.9739
9.3884 | -34.38 | -15.9832
9.3945 | -34.3804 | -15.9927
9.4006 | -34.3808 | -16.0023
9.4067 | -34.3812 | -16.0119
9.4128 | -34.3816 | -16.0214
9.4189 | -34.382 | -16.031
9.425 | -34.3824 | -16.0404
9.4311 | -34.3828 | -16.0499
9.4372 | -34.3832 | -16.0594
9.4433 | -34.3836 | -16.0689
9.4494 | -34.384 | -16.0781
9.4555 | -34.3844 | -16.0878
9.4616 | -34.3848 | -16.0975
9.4677 | -34.3852 | -16.107
9.4738 | -34.3856 | -16.1165
9.4799 | -34.386 | -16.126
9.486 | -34.3864 | -16.1354
9.4921 | -34.3868 | -16.145
9.4982 | -34.3872 | -16.1542
9.5043 | -34.3876 | -16.1636
9.5104 | -34.3879 | -16.1729
9.5165 | -34.3883 | -16.1823
9.5226 | -34.3887 | -16.1919
9.5288 | -34.3891 | -16.2011
9.5349 | -34.3895 | -16.2105
9.541 | -34.3899 | -16.2198
9.5471 | -34.3903 | -16.2292
9.5532 | -34.3907 | -16.2388
9.5593 | -34.3911 | -16.2484
9.5654 | -34.3915 | -16.2579
9.5715 | -34.3919 | -16.2673
9.5776 | -34.3923 | -16.2768
9.5837 | -34.3927 | -16.2863
9.5898 | -34.3931 | -16.2959
9.5959 | -34.3935 | -16.3054
9.602 | -34.3939 | -16.3147
9.6081 | -34.3943 | -16.3242
9.6142 | -34.3947 | -16.3336
9.6203 | -34.3951 | -16.3432
9.6264 | -34.3954 | -16.3528
9.6325 | -34.3958 | -16.3623
9.6386 | -34.3963 | -16.3716
9.6447 | -34.3967 | -16.3809
9.6508 | -34.397 | -16.3903
9.6569 | -34.3974 | -16.3998
9.663 | -34.3979 | -16.4093
9.6691 | -34.3982 | -16.4188
9.6753 | -34.3987 | -16.4282
9.6814 | -34.3991 | -16.4375
9.6875 | -34.3995 | -16.4467
9.6936 | -34.3999 | -16.4559
9.6997 | -34.4003 | -16.4656
9.7058 | -34.4007 | -16.4751
9.7119 | -34.4011 | -16.4845
9.718 | -34.4015 | -16.4941
9.7241 | -34.4019 | -16.5035
9.7302 | -34.4023 | -16.513
9.7363 | -34.4027 | -16.5223
9.7424 | -34.4031 | -16.5317
9.7485 | -34.4035 | -16.5412
9.7546 | -34.4038 | -16.5505
9.7607 | -34.4042 | -16.5602
9.7668 | -34.4047 | -16.5696
9.7729 | -34.4051 | -16.579
9.779 | -34.4055 | -16.588
9.7851 | -34.4059 | -16.5978
9.7912 | -34.4064 | -16.6071
9.7973 | -34.4068 | -16.6167
9.8034 | -34.4072 | -16.6259
9.8095 | -34.4076 | -16.6354
9.8157 | -34.408 | -16.6448
9.8218 | -34.4084 | -16.6543
9.8279 | -34.4088 | -16.6635
9.834 | -34.4092 | -16.6731
9.8401 | -34.4096 | -16.6824
9.8462 | -34.41 | -16.6917
9.8523 | -34.4104 | -16.701
9.8584 | -34.4108 | -16.7105
9.8645 | -34.4112 | -16.72
9.8706 | -34.4115 | -16.7294
9.8767 | -34.4119 | -16.7388
9.8828 | -34.4123 | -16.7482
9.8889 | -34.4127 | -16.7574
9.895 | -34.4131 | -16.7668
9.9011 | -34.4135 | -16.7762
9.9072 | -34.4139 | -16.7855
9.9133 | -34.4143 | -16.795
9.9194 | -34.4147 | -16.8043
9.9255 | -34.4151 | -16.8136
9.9316 | -34.4155 | -16.8229
9.9377 | -34.4159 | -16.8323
9.9438 | -34.4162 | -16.8418
9.9499 | -34.4166 | -16.8515
9.956 | -34.417 | -16.8609
9.9622 | -34.4174 | -16.8704
9.9683 | -34.4178 | -16.8798
9.9744 | -34.4182 | -16.8893
9.9805 | -34.4186 | -16.8986
9.9866 | -34.419 | -16.9081
9.9927 | -34.4194 | -16.9174
9.9988 | -34.4198 | -16.9269
10.0049 | -34.4202 | -16.9364
10.011 | -34.4206 | -16.9458
10.0171 | -34.421 | -16.9551
10.0232 | -34.4214 | -16.9644
10.0293 | -34.4218 | -16.9739
10.0354 | -34.4223 | -16.9833
10.0415 | -34.4227 | -16.9927
10.0476 | -34.4231 | -17.0022
10.0537 | -34.4234 | -17.0117
10.0598 | -34.4238 | -17.021
10.0659 | -34.4242 | -17.0304
10.072 | -34.4246 | -17.0397
10.0781 | -34.425 | -17.049
10.0842 | -34.4254 | -17.0583
10.0903 | -34.4258 | -17.0677
10.0964 | -34.4262 | -17.0771
10.1026 | -34.4266 | -17.0866
10.1087 | -34.427 | -17.0959
10.1148 | -34.4273 | -17.1054
10.1209 | -34.4277 | -17.1146
10.127 | -34.4281 | -17.1239
10.1331 | -34.4285 | -17.1331
10.1392 | -34.4289 | -17.1423
10.1453 | -34.4293 | -17.1516
10.1514 | -34.4297 | -17.161
10.1575 | -34.4301 | -17.1705
10.1636 | -34.4305 | -17.1796
10.1697 | -34.4309 | -17.1891
10.1758 | -34.4313 | -17.1983
10.1819 | -34.4317 | -17.2076
10.188 | -34.432 | -17.217
10.1941 | -34.4324 | -17.2263
10.2002 | -34.4328 | -17.2358
10.2063 | -34.4332 | -17.2451
10.2124 | -34.4336 | -17.2545
10.2185 | -34.434 | -17.2639
10.2246 | -34.4344 | -17.2732
10.2307 | -34.4348 | -17.2826
10.2368 | -34.4352 | -17.2914
10.2429 | -34.4356 | -17.3006
10.2491 | -34.4359 | -17.3101
10.2552 | -34.4363 | -17.3195
10.2613 | -34.4367 | -17.3288
10.2674 | -34.4371 | -17.3383
10.2735 | -34.4376 | -17.3478
10.2796 | -34.438 | -17.357
10.2857 | -34.4383 | -17.3664
10.2918 | -34.4387 | -17.3759
10.2979 | -34.4391 | -17.3853
10.304 | -34.4395 | -17.3945
10.3101 | -34.4399 | -17.4038
10.3162 | -34.4403 | -17.4133
10.3223 | -34.4407 | -17.4225
10.3284 | -34.4411 | -17.4317
10.3345 | -34.4415 | -17.441
10.3406 | -34.4419 | -17.4502
10.3467 | -34.4422 | -17.4595
10.3528 | -34.4426 | -17.4687
10.3589 | -34.443 | -17.478
10.365 | -34.4434 | -17.4874
10.3711 | -34.4438 | -17.4967
10.3772 | -34.4442 | -17.506
10.3833 | -34.4446 | -17.5153
10.3895 | -34.445 | -17.5246
10.3956 | -34.4454 | -17.5339
10.4017 | -34.4458 | -17.543
10.4078 | -34.4461 | -17.5523
10.4139 | -34.4465 | -17.5617
10.42 | -34.4469 | -17.5712
10.4261 | -34.4474 | -17.5805
10.4322 | -34.4477 | -17.5897
10.4383 | -34.4481 | -17.599
10.4444 | -34.4485 | -17.6082
10.4505 | -34.4489 | -17.6176
10.4566 | -34.4492 | -17.6267
10.4627 | -34.4496 | -17.636
10.4688 | -34.45 | -17.6451
10.4749 | -34.4504 | -17.6544
10.481 | -34.4508 | -17.6636
10.4871 | -34.4512 | -17.6728
10.4932 | -34.4516 | -17.682
10.4993 | -34.452 | -17.6914
10.5054 | -34.4524 | -17.7007
10.5115 | -34.4528 | -17.7099
10.5176 | -34.4532 | -17.7191
10.5237 | -34.4536 | -17.7285
10.5298 | -34.454 | -17.7377
10.536 | -34.4544 | -17.7471
10.5421 | -34.4548 | -17.7561
10.5482 | -34.4551 | -17.7655
10.5543 | -34.4556 | -17.7746
10.5604 | -34.4559 | -17.7836
10.5665 | -34.4563 | -17.7927
10.5726 | -34.4567 | -17.802
10.5787 | -34.4571 | -17.8112
10.5848 | -34.4575 | -17.8203
10.5909 | -34.4579 | -17.8297
10.597 | -34.4583 | -17.8389
10.6031 | -34.4587 | -17.8484
10.6092 | -34.459 | -17.8577
10.6153 | -34.4594 | -17.8671
10.6214 | -34.4598 | -17.8762
10.6275 | -34.4602 | -17.8855
10.6336 | -34.4606 | -17.8948
10.6397 | -34.461 | -17.9039
10.6458 | -34.4614 | -17.9131
10.6519 | -34.4618 | -17.9225
10.658 | -34.4622 | -17.9317
10.6641 | -34.4626 | -17.9411
10.6702 | -34.463 | -17.9505
10.6764 | -34.4634 | -17.9598
10.6825 | -34.4638 | -17.9689
10.6886 | -34.4642 | -17.978
10.6947 | -34.4646 | -17.9872
10.7008 | -34.465 | -17.9964
10.7069 | -34.4654 | -18.0057
10.713 | -34.4658 | -18.015
10.7191 | -34.4662 | -18.0242
10.7252 | -34.4665 | -18.0332
10.7313 | -34.4669 | -18.0424
10.7374 | -34.4673 | -18.0517
10.7435 | -34.4677 | -18.061
10.7496 | -34.4681 | -18.0699
10.7557 | -34.4684 | -18.0792
10.7618 | -34.4689 | -18.0882
10.7679 | -34.4693 | -18.0974
10.774 | -34.4697 | -18.1066
10.7801 | -34.4701 | -18.1158
10.7862 | -34.4705 | -18.1248
10.7923 | -34.4708 | -18.1341
10.7984 | -34.4713 | -18.1432
10.8045 | -34.4717 | -18.1523
10.8106 | -34.472 | -18.1616
10.8168 | -34.4724 | -18.1708
10.8229 | -34.4728 | -18.18
10.829 | -34.4732 | -18.1892
10.8351 | -34.4736 | -18.1982
10.8412 | -34.474 | -18.2073
10.8473 | -34.4744 | -18.2163
10.8534 | -34.4748 | -18.2254
10.8595 | -34.4753 | -18.2346
10.8656 | -34.4757 | -18.2436
10.8717 | -34.4761 | -18.2527
10.8778 | -34.4765 | -18.2619
10.8839 | -34.4769 | -18.2711
10.89 | -34.4773 | -18.2803
10.8961 | -34.4777 | -18.2894
10.9022 | -34.4781 | -18.2986
10.9083 | -34.4785 | -18.3077
10.9144 | -34.4789 | -18.3168
10.9205 | -34.4792 | -18.326
10.9266 | -34.4796 | -18.3351
10.9327 | -34.4801 | -18.3441
10.9388 | -34.4805 | -18.3529
10.9449 | -34.4809 | -18.362
10.951 | -34.4813 | -18.3711
10.9571 | -34.4817 | -18.3803
10.9633 | -34.4821 | -18.3896
10.9694 | -34.4825 | -18.399
10.9755 | -34.4829 | -18.4081
10.9816 | -34.4833 | -18.4174
10.9877 | -34.4837 | -18.4263
10.9938 | -34.4841 | -18.4357
10.9999 | -34.4846 | -18.4447
11.006 | -34.485 | -18.4535
11.0121 | -34.4854 | -18.4627
11.0182 | -34.4858 | -18.4719
11.0243 | -34.4862 | -18.481
11.0304 | -34.4866 | -18.4901
11.0365 | -34.487 | -18.4991
11.0426 | -34.4874 | -18.5081
11.0487 | -34.4879 | -18.5173
11.0548 | -34.4883 | -18.5262
11.0609 | -34.4887 | -18.5354
11.067 | -34.4891 | -18.5445
11.0731 | -34.4895 | -18.5535
11.0792 | -34.49 | -18.5625
11.0853 | -34.4904 | -18.5716
11.0914 | -34.4908 | -18.5808
11.0975 | -34.4912 | -18.5899
11.1037 | -34.4916 | -18.5989
11.1098 | -34.492 | -18.6082
11.1159 | -34.4924 | -18.617
11.122 | -34.4928 | -18.6259
11.1281 | -34.4932 | -18.6349
11.1342 | -34.4936 | -18.644
11.1403 | -34.494 | -18.6532
11.1464 | -34.4944 | -18.6623
11.1525 | -34.4948 | -18.6712
11.1586 | -34.4952 | -18.6804
11.1647 | -34.4956 | -18.6894
11.1708 | -34.496 | -18.6985
11.1769 | -34.4964 | -18.7076
11.183 | -34.4969 | -18.7166
11.1891 | -34.4973 | -18.7255
11.1952 | -34.4978 | -18.7345
11.2013 | -34.4982 | -18.7436
11.2074 | -34.4986 | -18.7527
11.2135 | -34.499 | -18.7619
11.2196 | -34.4994 | -18.7709
11.2257 | -34.4999 | -18.78
11.2318 | -34.5003 | -18.789
11.2379 | -34.5008 | -18.7983
11.244 | -34.5012 | -18.8075
11.2502 | -34.5016 | -18.8165
11.2563 | -34.5021 | -18.8256
11.2624 | -34.5025 | -18.8345
11.2685 | -34.5029 | -18.8434
11.2746 | -34.5033 | -18.8523
11.2807 | -34.5038 | -18.8616
11.2868 | -34.5042 | -18.8707
11.2929 | -34.5046 | -18.8794
11.299 | -34.5051 | -18.8885
11.3051 | -34.5055 | -18.8977
11.3112 | -34.5059 | -18.9068
11.3173 | -34.5063 | -18.9157
11.3234 | -34.5068 | -18.9245
11.3295 | -34.5072 | -18.9337
11.3356 | -34.5076 | -18.9428
11.3417 | -34.508 | -18.9517
11.3478 | -34.5085 | -18.961
11.3539 | -34.5089 | -18.9701
11.36 | -34.5093 | -18.979
11.3661 | -34.5097 | -18.9878
11.3722 | -34.5101 | -18.9965
11.3783 | -34.5106 | -19.0056
11.3844 | -34.511 | -19.0146
11.3906 | -34.5114 | -19.0236
11.3967 | -34.5118 | -19.0325
11.4028 | -34.5123 | -19.0418
11.4089 | -34.5127 | -19.0506
11.415 | -34.5131 | -19.0595
11.4211 | -34.5135 | -19.0684
11.4272 | -34.5139 | -19.0774
11.4333 | -34.5144 | -19.0861
11.4394 | -34.5148 | -19.0951
11.4455 | -34.5152 | -19.104
11.4516 | -34.5156 | -19.113
11.4577 | -34.516 | -19.1219
11.4638 | -34.5165 | -19.1307
11.4699 | -34.5169 | -19.1399
11.476 | -34.5173 | -19.1491
11.4821 | -34.5178 | -19.1581
11.4882 | -34.5182 | -19.1672
11.4943 | -34.5187 | -19.1764
11.5004 | -34.5191 | -19.1851
11.5065 | -34.5196 | -19.194
11.5126 | -34.52 | -19.2029
11.5187 | -34.5204 | -19.2119
11.5248 | -34.5208 | -19.2212
11.5309 | -34.5213 | -19.2301
11.5371 | -34.5217 | -19.2389
11.5432 | -34.5222 | -19.2479
11.5493 | -34.5226 | -19.257
11.5554 | -34.523 | -19.266
11.5615 | -34.5235 | -19.2751
11.5676 | -34.5239 | -19.284
11.5737 | -34.5244 | -19.2929
11.5798 | -34.5248 | -19.3017
11.5859 | -34.5253 | -19.3106
11.592 | -34.5257 | -19.3195
11.5981 | -34.5261 | -19.3284
11.6042 | -34.5265 | -19.3373
11.6103 | -34.527 | -19.3463
11.6164 | -34.5274 | -19.3551
11.6225 | -34.5278 | -19.364
11.6286 | -34.5282 | -19.3729
11.6347 | -34.5287 | -19.382
11.6408 | -34.5291 | -19.3908
11.6469 | -34.5296 | -19.3996
11.653 | -34.53 | -19.4085
11.6591 | -34.5304 | -19.4174
11.6652 | -34.5308 | -19.4262
11.6713 | -34.5313 | -19.435
11.6775 | -34.5317 | -19.4438
11.6836 | -34.5322 | -19.4527
11.6897 | -34.5326 | -19.4617
11.6958 | -34.533 | -19.4706
11.7019 | -34.5335 | -19.4797
11.708 | -34.5339 | -19.4885
11.7141 | -34.5344 | -19.4973
11.7202 | -34.5349 | -19.5062
11.7263 | -34.5353 | -19.5152
11.7324 | -34.5357 | -19.5244
11.7385 | -34.5361 | -19.5334
11.7446 | -34.5365 | -19.5422
11.7507 | -34.537 | -19.551
11.7568 | -34.5374 | -19.56
11.7629 | -34.5379 | -19.569
11.769 | -34.5383 | -19.5777
11.7751 | -34.5387 | -19.5866
11.7812 | -34.5392 | -19.5953
11.7873 | -34.5397 | -19.6042
11.7934 | -34.5401 | -19.6132
11.7995 | -34.5406 | -19.622
11.8056 | -34.541 | -19.6309
11.8117 | -34.5415 | -19.6398
11.8178 | -34.5419 | -19.6485
11.824 | -34.5424 | -19.6573
11.8301 | -34.5428 | -19.6661
11.8362 | -34.5433 | -19.6749
11.8423 | -34.5437 | -19.6837
11.8484 | -34.5441 | -19.6926
11.8545 | -34.5446 | -19.7014
11.8606 | -34.5451 | -19.7101
11.8667 | -34.5455 | -19.7191
11.8728 | -34.546 | -19.7278
11.8789 | -34.5464 | -19.7366
11.885 | -34.5468 | -19.7454
11.8911 | -34.5473 | -19.7542
11.8972 | -34.5478 | -19.763
11.9033 | -34.5482 | -19.7718
11.9094 | -34.5486 | -19.7807
11.9155 | -34.5491 | -19.7895
11.9216 | -34.5496 | -19.7983
11.9277 | -34.55 | -19.8071
11.9338 | -34.5504 | -19.8159
11.9399 | -34.5509 | -19.8249
11.946 | -34.5514 | -19.8337
11.9521 | -34.5518 | -19.8423
11.9582 | -34.5523 | -19.8511
11.9644 | -34.5527 | -19.8599
11.9705 | -34.5532 | -19.8686
11.9766 | -34.5537 | -19.8773
11.9827 | -34.5541 | -19.886
11.9888 | -34.5546 | -19.8948
11.9949 | -34.5551 | -19.9034
12.001 | -34.5555 | -19.9122
12.0071 | -34.556 | -19.9209
12.0132 | -34.5564 | -19.9296
12.0193 | -34.5569 | -19.9384
12.0254 | -34.5573 | -19.9471
12.0315 | -34.5577 | -19.9557
12.0376 | -34.5582 | -19.9644
12.0437 | -34.5586 | -19.9731
12.0498 | -34.5591 | -19.982
12.0559 | -34.5595 | -19.9906
12.062 | -34.56 | -19.9991
12.0681 | -34.5605 | -20.0079
12.0742 | -34.561 | -20.0165
12.0803 | -34.5614 | -20.0255
12.0864 | -34.5618 | -20.0343
12.0925 | -34.5623 | -20.043
12.0986 | -34.5628 | -20.0517
12.1047 | -34.5632 | -20.0605
12.1109 | -34.5637 | -20.0691
12.117 | -34.5642 | -20.0778
12.1231 | -34.5646 | -20.0865
12.1292 | -34.5651 | -20.0952
12.1353 | -34.5655 | -20.1039
12.1414 | -34.566 | -20.1127
12.1475 | -34.5664 | -20.1216
12.1536 | -34.5669 | -20.1303
12.1597 | -34.5673 | -20.139
12.1658 | -34.5678 | -20.1477
12.1719 | -34.5683 | -20.1562
12.178 | -34.5687 | -20.165
12.1841 | -34.5692 | -20.1738
12.1902 | -34.5696 | -20.1826
12.1963 | -34.5701 | -20.1914
12.2024 | -34.5706 | -20.1998
12.2085 | -34.571 | -20.2086
12.2146 | -34.5715 | -20.2174
12.2207 | -34.572 | -20.2261
12.2268 | -34.5724 | -20.2348
12.2329 | -34.5729 | -20.2434
12.239 | -34.5734 | -20.2518
12.2451 | -34.5738 | -20.2605
12.2513 | -34.5743 | -20.2694
12.2574 | -34.5747 | -20.2781
12.2635 | -34.5752 | -20.2868
12.2696 | -34.5757 | -20.2954
12.2757 | -34.5762 | -20.3038
12.2818 | -34.5766 | -20.3126
12.2879 | -34.5771 | -20.3213
12.294 | -34.5776 | -20.3298
12.3001 | -34.578 | -20.3385
12.3062 | -34.5785 | -20.3472
12.3123 | -34.579 | -20.3555
12.3184 | -34.5795 | -20.3641
12.3245 | -34.5799 | -20.3727
12.3306 | -34.5804 | -20.3814
12.3367 | -34.5808 | -20.3899
12.3428 | -34.5813 | -20.3986
12.3489 | -34.5818 | -20.407
12.355 | -34.5822 | -20.4158
12.3611 | -34.5827 | -20.4247
12.3672 | -34.5831 | -20.4332
12.3733 | -34.5836 | -20.4419
12.3794 | -34.5841 | -20.4504
12.3855 | -34.5845 | -20.4589
12.3916 | -34.585 | -20.4677
12.3978 | -34.5855 | -20.4765
12.4039 | -34.5859 | -20.4851
12.41 | -34.5864 | -20.4938
12.4161 | -34.5868 | -20.5021
12.4222 | -34.5873 | -20.5107
12.4283 | -34.5878 | -20.5191
12.4344 | -34.5882 | -20.5279
12.4405 | -34.5887 | -20.5365
12.4466 | -34.5891 | -20.5451
12.4527 | -34.5896 | -20.5537
12.4588 | -34.59 | -20.5623
12.4649 | -34.5905 | -20.571
12.471 | -34.591 | -20.5797
12.4771 | -34.5914 | -20.5883
12.4832 | -34.5919 | -20.5968
12.4893 | -34.5924 | -20.6052
12.4954 | -34.5928 | -20.6137
12.5015 | -34.5933 | -20.6223
12.5076 | -34.5938 | -20.631
12.5137 | -34.5943 | -20.6397
12.5198 | -34.5947 | -20.6482
12.5259 | -34.5952 | -20.6568
12.532 | -34.5956 | -20.6654
12.5382 | -34.5961 | -20.6741
12.5443 | -34.5965 | -20.6826
12.5504 | -34.597 | -20.6913
12.5565 | -34.5975 | -20.6998
12.5626 | -34.598 | -20.7084
12.5687 | -34.5984 | -20.717
12.5748 | -34.5989 | -20.7257
12.5809 | -34.5993 | -20.7342
12.587 | -34.5998 | -20.7429
12.5931 | -34.6003 | -20.7513
12.5992 | -34.6008 | -20.7598
12.6053 | -34.6012 | -20.7684
12.6114 | -34.6017 | -20.7768
12.6175 | -34.6022 | -20.7855
12.6236 | -34.6026 | -20.7941
12.6297 | -34.6031 | -20.8027
12.6358 | -34.6036 | -20.8111
12.6419 | -34.604 | -20.8196
12.648 | -34.6045 | -20.8282
12.6541 | -34.605 | -20.8368
12.6602 | -34.6055 | -20.8448
12.6663 | -34.6059 | -20.8533
12.6724 | -34.6064 | -20.8617
12.6785 | -34.6069 | -20.8703
12.6847 | -34.6073 | -20.8787
12.6908 | -34.6078 | -20.8872
12.6969 | -34.6083 | -20.8956
12.703 | -34.6087 | -20.9042
12.7091 | -34.6092 | -20.9127
12.7152 | -34.6096 | -20.9211
12.7213 | -34.6101 | -20.9298
12.7274 | -34.6106 | -20.9383
12.7335 | -34.611 | -20.9466
12.7396 | -34.6115 | -20.9549
12.7457 | -34.6119 | -20.9636
12.7518 | -34.6124 | -20.972
12.7579 | -34.6129 | -20.9806
12.764 | -34.6133 | -20.9894
12.7701 | -34.6138 | -20.9981
12.7762 | -34.6142 | -21.0065
12.7823 | -34.6147 | -21.015
12.7884 | -34.6152 | -21.0235
12.7945 | -34.6157 | -21.0323
12.8006 | -34.6161 | -21.0409
12.8067 | -34.6166 | -21.0494
12.8128 | -34.617 | -21.0578
12.8189 | -34.6175 | -21.0662
12.8251 | -34.618 | -21.0747
12.8312 | -34.6184 | -21.0833
12.8373 | -34.6189 | -21.0917
12.8434 | -34.6193 | -21.1003
12.8495 | -34.6198 | -21.1086
12.8556 | -34.6202 | -21.117
12.8617 | -34.6207 | -21.1256
12.8678 | -34.6212 | -21.1343
12.8739 | -34.6216 | -21.1428
12.88 | -34.6221 | -21.1512
12.8861 | -34.6225 | -21.1597
12.8922 | -34.623 | -21.1681
12.8983 | -34.6235 | -21.1766
12.9044 | -34.6239 | -21.1849
12.9105 | -34.6244 | -21.1933
12.9166 | -34.6249 | -21.2018
12.9227 | -34.6253 | -21.2102
12.9288 | -34.6258 | -21.2187
12.9349 | -34.6262 | -21.2272
12.941 | -34.6267 | -21.2357
12.9471 | -34.6272 | -21.2442
12.9532 | -34.6277 | -21.2526
12.9593 | -34.6281 | -21.2613
12.9654 | -34.6286 | -21.2699
12.9716 | -34.6291 | -21.2786
12.9777 | -34.6296 | -21.2872
12.9838 | -34.6301 | -21.2956
12.9899 | -34.6305 | -21.3041
12.996 | -34.631 | -21.3126
13.0021 | -34.6315 | -21.3211
13.0082 | -34.6319 | -21.3295
13.0143 | -34.6324 | -21.3378
13.0204 | -34.6329 | -21.3461
13.0265 | -34.6334 | -21.3546
13.0326 | -34.6338 | -21.363
13.0387 | -34.6343 | -21.3715
13.0448 | -34.6347 | -21.3799
13.0509 | -34.6352 | -21.3884
13.057 | -34.6357 | -21.3968
13.0631 | -34.6362 | -21.4051
13.0692 | -34.6366 | -21.4136
13.0753 | -34.6371 | -21.4222
13.0814 | -34.6376 | -21.4307
13.0875 | -34.6381 | -21.439
13.0936 | -34.6385 | -21.4476
13.0997 | -34.639 | -21.4561
13.1058 | -34.6394 | -21.4644
13.112 | -34.6398 | -21.4727
13.1181 | -34.6403 | -21.4811
13.1242 | -34.6408 | -21.4897
13.1303 | -34.6412 | -21.4979
13.1364 | -34.6417 | -21.5065
13.1425 | -34.6422 | -21.515
13.1486 | -34.6426 | -21.5234
13.1547 | -34.6431 | -21.532
13.1608 | -34.6435 | -21.5404
13.1669 | -34.644 | -21.5488
13.173 | -34.6444 | -21.5571
13.1791 | -34.6449 | -21.5654
13.1852 | -34.6453 | -21.5737
13.1913 | -34.6458 | -21.5819
13.1974 | -34.6462 | -21.5906
13.2035 | -34.6467 | -21.5991
13.2096 | -34.6471 | -21.6073
13.2157 | -34.6476 | -21.6156
13.2218 | -34.6481 | -21.624
13.2279 | -34.6485 | -21.6323
13.234 | -34.649 | -21.6406
13.2401 | -34.6495 | -21.649
13.2462 | -34.6499 | -21.6575
13.2524 | -34.6504 | -21.6659
13.2585 | -34.6508 | -21.6743
13.2646 | -34.6513 | -21.6828
13.2707 | -34.6518 | -21.6912
13.2768 | -34.6522 | -21.6997
13.2829 | -34.6527 | -21.7082
13.289 | -34.6532 | -21.7164
13.2951 | -34.6536 | -21.7245
13.3012 | -34.6541 | -21.733
13.3073 | -34.6546 | -21.7414
13.3134 | -34.655 | -21.7497
13.3195 | -34.6555 | -21.7581
13.3256 | -34.6559 | -21.7663
13.3317 | -34.6564 | -21.7747
13.3378 | -34.6568 | -21.7832
13.3439 | -34.6573 | -21.7917
13.35 | -34.6577 | -21.8002
13.3561 | -34.6582 | -21.8085
13.3622 | -34.6586 | -21.8168
13.3683 | -34.6591 | -21.825
13.3744 | -34.6595 | -21.8334
13.3805 | -34.66 | -21.8421
13.3866 | -34.6604 | -21.8504
13.3927 | -34.6608 | -21.8588
13.3989 | -34.6613 | -21.8674
13.405 | -34.6617 | -21.8758
13.4111 | -34.6621 | -21.884
13.4172 | -34.6626 | -21.8925
13.4233 | -34.663 | -21.901
13.4294 | -34.6635 | -21.9096
13.4355 | -34.6639 | -21.9177
13.4416 | -34.6644 | -21.926
13.4477 | -34.6649 | -21.9345
13.4538 | -34.6653 | -21.9427
13.4599 | -34.6658 | -21.9512
13.466 | -34.6662 | -21.9595
13.4721 | -34.6666 | -21.9677
13.4782 | -34.6671 | -21.9761
13.4843 | -34.6676 | -21.9846
13.4904 | -34.668 | -21.9928
13.4965 | -34.6685 | -22.0014
13.5026 | -34.6689 | -22.0098
13.5087 | -34.6694 | -22.0179
13.5148 | -34.6699 | -22.0264
13.5209 | -34.6703 | -22.035
13.527 | -34.6707 | -22.0434
13.5331 | -34.6712 | -22.0519
13.5393 | -34.6717 | -22.0601
13.5454 | -34.6721 | -22.0685
13.5515 | -34.6726 | -22.0771
13.5576 | -34.673 | -22.0855
13.5637 | -34.6735 | -22.0939
13.5698 | -34.6739 | -22.1023
13.5759 | -34.6744 | -22.1103
13.582 | -34.6748 | -22.1186
13.5881 | -34.6753 | -22.1269
13.5942 | -34.6758 | -22.1354
13.6003 | -34.6762 | -22.1436
13.6064 | -34.6767 | -22.1522
13.6125 | -34.6771 | -22.1607
13.6186 | -34.6776 | -22.169
13.6247 | -34.6781 | -22.1774
13.6308 | -34.6785 | -22.1858
13.6369 | -34.6789 | -22.1943
13.643 | -34.6794 | -22.2028
13.6491 | -34.6798 | -22.2111
13.6552 | -34.6803 | -22.2193
13.6613 | -34.6807 | -22.2277
13.6674 | -34.6812 | -22.2361
13.6735 | -34.6816 | -22.2444
13.6796 | -34.6821 | -22.2528
13.6858 | -34.6825 | -22.2609
13.6919 | -34.683 | -22.2694
13.698 | -34.6834 | -22.2779
13.7041 | -34.6839 | -22.2863
13.7102 | -34.6843 | -22.2948
13.7163 | -34.6848 | -22.303
13.7224 | -34.6852 | -22.3113
13.7285 | -34.6857 | -22.3197
13.7346 | -34.6861 | -22.3281
13.7407 | -34.6866 | -22.3364
13.7468 | -34.6871 | -22.3448
13.7529 | -34.6875 | -22.353
13.759 | -34.6879 | -22.3614
13.7651 | -34.6884 | -22.3698
13.7712 | -34.6889 | -22.3782
13.7773 | -34.6893 | -22.3867
13.7834 | -34.6897 | -22.3952
13.7895 | -34.6902 | -22.4035
13.7956 | -34.6907 | -22.412
13.8017 | -34.6911 | -22.4202
13.8078 | -34.6915 | -22.4288
13.8139 | -34.692 | -22.4369
13.82 | -34.6924 | -22.4451
13.8262 | -34.6929 | -22.4535
13.8323 | -34.6933 | -22.4621
13.8384 | -34.6938 | -22.4706
13.8445 | -34.6943 | -22.479
13.8506 | -34.6947 | -22.4874
13.8567 | -34.6951 | -22.4958
13.8628 | -34.6956 | -22.5041
13.8689 | -34.696 | -22.5126
13.875 | -34.6965 | -22.521
13.8811 | -34.6969 | -22.5294
13.8872 | -34.6974 | -22.5378
13.8933 | -34.6978 | -22.5461
13.8994 | -34.6982 | -22.5545
13.9055 | -34.6987 | -22.5631
13.9116 | -34.6991 | -22.5713
13.9177 | -34.6995 | -22.5797
13.9238 | -34.7 | -22.588
13.9299 | -34.7004 | -22.5963
13.936 | -34.7009 | -22.6047
13.9421 | -34.7013 | -22.6132
13.9482 | -34.7017 | -22.6214
13.9543 | -34.7021 | -22.6299
13.9604 | -34.7026 | -22.6383
13.9665 | -34.703 | -22.6466
13.9727 | -34.7035 | -22.6552
13.9788 | -34.7039 | -22.6634
13.9849 | -34.7044 | -22.6719
13.991 | -34.7049 | -22.6802
13.9971 | -34.7053 | -22.6883
14.0032 | -34.7057 | -22.6966
14.0093 | -34.7062 | -22.705
14.0154 | -34.7066 | -22.7132
14.0215 | -34.707 | -22.7214
14.0276 | -34.7075 | -22.7298
14.0337 | -34.708 | -22.7381
14.0398 | -34.7085 | -22.7464
14.0459 | -34.7089 | -22.7547
14.052 | -34.7094 | -22.7631
14.0581 | -34.7098 | -22.7716
14.0642 | -34.7102 | -22.7799
14.0703 | -34.7107 | -22.7883
14.0764 | -34.7111 | -22.7965
14.0825 | -34.7116 | -22.805
14.0886 | -34.712 | -22.8133
14.0947 | -34.7125 | -22.8216
14.1008 | -34.7129 | -22.8299
14.1069 | -34.7134 | -22.8381
14.1131 | -34.7138 | -22.8465
14.1192 | -34.7143 | -22.8548
14.1253 | -34.7147 | -22.8631
14.1314 | -34.7152 | -22.8713
14.1375 | -34.7156 | -22.8796
14.1436 | -34.7161 | -22.8879
14.1497 | -34.7166 | -22.8962
14.1558 | -34.717 | -22.9045
14.1619 | -34.7174 | -22.9127
14.168 | -34.7179 | -22.921
14.1741 | -34.7183 | -22.9292
14.1802 | -34.7188 | -22.9375
14.1863 | -34.7192 | -22.946
14.1924 | -34.7197 | -22.9542
14.1985 | -34.7202 | -22.9624
14.2046 | -34.7206 | -22.9708
14.2107 | -34.7211 | -22.9788
14.2168 | -34.7215 | -22.9872
14.2229 | -34.7219 | -22.9954
14.229 | -34.7224 | -23.0039
14.2351 | -34.7228 | -23.0122
14.2412 | -34.7233 | -23.0205
14.2473 | -34.7238 | -23.0287
14.2534 | -34.7242 | -23.0368
14.2596 | -34.7246 | -23.0449
14.2657 | -34.725 | -23.0531
14.2718 | -34.7255 | -23.0615
14.2779 | -34.726 | -23.0699
14.284 | -34.7264 | -23.0783
14.2901 | -34.7269 | -23.0866
14.2962 | -34.7273 | -23.0949
14.3023 | -34.7277 | -23.1033
14.3084 | -34.7282 | -23.1116
14.3145 | -34.7287 | -23.1199
14.3206 | -34.7291 | -23.1284
14.3267 | -34.7296 | -23.1367
14.3328 | -34.73 | -23.145
14.3389 | -34.7305 | -23.1532
14.345 | -34.7309 | -23.1614
14.3511 | -34.7313 | -23.1697
14.3572 | -34.7318 | -23.178
14.3633 | -34.7322 | -23.1864
14.3694 | -34.7327 | -23.1947
14.3755 | -34.7332 | -23.203
14.3816 | -34.7336 | -23.2113
14.3877 | -34.7341 | -23.2195
14.3938 | -34.7346 | -23.228
14.4 | -34.735 | -23.2363
14.4061 | -34.7355 | -23.2445
14.4122 | -34.7359 | -23.2529
14.4183 | -34.7363 | -23.2611
14.4244 | -34.7368 | -23.2697
14.4305 | -34.7373 | -23.2779
14.4366 | -34.7377 | -23.2861
14.4427 | -34.7382 | -23.2946
14.4488 | -34.7386 | -23.303
14.4549 | -34.7391 | -23.3112
14.461 | -34.7396 | -23.3193
14.4671 | -34.74 | -23.3277
14.4732 | -34.7405 | -23.3361
14.4793 | -34.7409 | -23.3444
14.4854 | -34.7414 | -23.3526
14.4915 | -34.7419 | -23.3606
14.4976 | -34.7423 | -23.3689
14.5037 | -34.7428 | -23.3771
14.5098 | -34.7432 | -23.3854
14.5159 | -34.7437 | -23.3938
14.522 | -34.7441 | -23.4019
14.5281 | -34.7446 | -23.4101
14.5342 | -34.745 | -23.4184
14.5403 | -34.7455 | -23.4267
14.5465 | -34.7459 | -23.435
14.5526 | -34.7464 | -23.4435
14.5587 | -34.7469 | -23.4521
14.5648 | -34.7473 | -23.4603
14.5709 | -34.7478 | -23.4684
14.577 | -34.7482 | -23.4767
14.5831 | -34.7487 | -23.485
14.5892 | -34.7491 | -23.4932
14.5953 | -34.7496 | -23.5013
14.6014 | -34.75 | -23.5096
14.6075 | -34.7504 | -23.5178
14.6136 | -34.7509 | -23.5262
14.6197 | -34.7513 | -23.5346
14.6258 | -34.7518 | -23.543
14.6319 | -34.7523 | -23.5512
14.638 | -34.7527 | -23.5595
14.6441 | -34.7532 | -23.5678
14.6502 | -34.7536 | -23.5762
14.6563 | -34.7541 | -23.5844
14.6624 | -34.7546 | -23.5928
14.6685 | -34.7551 | -23.6008
14.6746 | -34.7555 | -23.6091
14.6807 | -34.756 | -23.6172
14.6869 | -34.7565 | -23.6253
14.693 | -34.757 | -23.6335
14.6991 | -34.7574 | -23.6417
14.7052 | -34.7579 | -23.6499
14.7113 | -34.7584 | -23.6582
14.7174 | -34.7588 | -23.6665
14.7235 | -34.7593 | -23.6747
14.7296 | -34.7598 | -23.6829
14.7357 | -34.7603 | -23.6911
14.7418 | -34.7607 | -23.6994
14.7479 | -34.7612 | -23.7076
14.754 | -34.7616 | -23.7159
14.7601 | -34.7621 | -23.7242
14.7662 | -34.7626 | -23.7324
14.7723 | -34.763 | -23.7404
14.7784 | -34.7635 | -23.7487
14.7845 | -34.7639 | -23.7571
14.7906 | -34.7644 | -23.7653
14.7967 | -34.7649 | -23.7736
14.8028 | -34.7653 | -23.7818
14.8089 | -34.7658 | -23.79
14.815 | -34.7663 | -23.7982
14.8211 | -34.7667 | -23.8063
14.8272 | -34.7672 | -23.8148
14.8334 | -34.7676 | -23.8229
14.8395 | -34.7681 | -23.831
14.8456 | -34.7686 | -23.8392
14.8517 | -34.7691 | -23.8475
14.8578 | -34.7695 | -23.8557
14.8639 | -34.77 | -23.864
14.87 | -34.7705 | -23.8722
14.8761 | -34.771 | -23.8804
14.8822 | -34.7714 | -23.8888
14.8883 | -34.7719 | -23.8971
14.8944 | -34.7723 | -23.9052
14.9005 | -34.7728 | -23.9136
14.9066 | -34.7732 | -23.9218
14.9127 | -34.7737 | -23.9299
14.9188 | -34.7742 | -23.9381
14.9249 | -34.7747 | -23.9463
14.931 | -34.7751 | -23.9545
14.9371 | -34.7756 | -23.9627
14.9432 | -34.776 | -23.9709
14.9493 | -34.7766 | -23.9789
14.9554 | -34.777 | -23.9871
14.9615 | -34.7775 | -23.9953
14.9676 | -34.7779 | -24.0037
14.9738 | -34.7784 | -24.0117
14.9799 | -34.7788 | -24.0201
14.986 | -34.7793 | -24.0282
14.9921 | -34.7798 | -24.0364
14.9982 | -34.7802 | -24.0446
15.0043 | -34.7807 | -24.0528
15.0104 | -34.7812 | -24.061
15.0165 | -34.7817 | -24.0694
15.0226 | -34.7822 | -24.0772
15.0287 | -34.7826 | -24.0855
15.0348 | -34.7831 | -24.0937
15.0409 | -34.7836 | -24.1018
15.047 | -34.7841 | -24.11
15.0531 | -34.7846 | -24.1181
15.0592 | -34.785 | -24.1264
15.0653 | -34.7855 | -24.1346
15.0714 | -34.786 | -24.1426
15.0775 | -34.7864 | -24.1507
15.0836 | -34.7869 | -24.1588
15.0897 | -34.7874 | -24.1668
15.0958 | -34.7879 | -24.1751
15.1019 | -34.7884 | -24.1832
15.108 | -34.7889 | -24.1912
15.1141 | -34.7893 | -24.1993
15.1203 | -34.7898 | -24.2076
15.1264 | -34.7903 | -24.2157
15.1325 | -34.7908 | -24.2239
15.1386 | -34.7912 | -24.2319
15.1447 | -34.7917 | -24.24
15.1508 | -34.7922 | -24.2482
15.1569 | -34.7926 | -24.2562
15.163 | -34.7931 | -24.2642
15.1691 | -34.7936 | -24.2723
15.1752 | -34.7941 | -24.2805
15.1813 | -34.7945 | -24.2887
15.1874 | -34.795 | -24.2968
15.1935 | -34.7955 | -24.3047
15.1996 | -34.796 | -24.3127
15.2057 | -34.7964 | -24.3209
15.2118 | -34.7969 | -24.329
15.2179 | -34.7974 | -24.3372
15.224 | -34.7978 | -24.3451
15.2301 | -34.7983 | -24.353
15.2362 | -34.7988 | -24.3612
15.2423 | -34.7993 | -24.3695
15.2484 | -34.7998 | -24.3776
15.2545 | -34.8003 | -24.3859
15.2607 | -34.8008 | -24.3941
15.2668 | -34.8013 | -24.4019
15.2729 | -34.8018 | -24.4099
15.279 | -34.8023 | -24.4184
15.2851 | -34.8028 | -24.4265
15.2912 | -34.8032 | -24.4347
15.2973 | -34.8037 | -24.4429
15.3034 | -34.8042 | -24.4511
15.3095 | -34.8047 | -24.4592
15.3156 | -34.8052 | -24.4673
15.3217 | -34.8057 | -24.4754
15.3278 | -34.8061 | -24.4836
15.3339 | -34.8066 | -24.4914
15.34 | -34.8071 | -24.4993
15.3461 | -34.8076 | -24.5073
15.3522 | -34.8081 | -24.5152
15.3583 | -34.8085 | -24.5234
15.3644 | -34.809 | -24.5313
15.3705 | -34.8095 | -24.5389
15.3766 | -34.81 | -24.5471
15.3827 | -34.8105 | -24.555
15.3888 | -34.8109 | -24.5632
15.3949 | -34.8114 | -24.5713
15.401 | -34.8119 | -24.5794
15.4072 | -34.8124 | -24.5871
15.4133 | -34.8129 | -24.5952
15.4194 | -34.8134 | -24.6033
15.4255 | -34.8138 | -24.6114
15.4316 | -34.8143 | -24.6195
15.4377 | -34.8148 | -24.6278
15.4438 | -34.8153 | -24.6357
15.4499 | -34.8158 | -24.6437
15.456 | -34.8163 | -24.6518
15.4621 | -34.8168 | -24.6598
15.4682 | -34.8173 | -24.6679
15.4743 | -34.8178 | -24.6756
15.4804 | -34.8182 | -24.6837
15.4865 | -34.8187 | -24.6917
15.4926 | -34.8192 | -24.6998
15.4987 | -34.8197 | -24.7078
15.5048 | -34.8202 | -24.7158
15.5109 | -34.8206 | -24.7237
15.517 | -34.8212 | -24.7318
15.5231 | -34.8216 | -24.7398
15.5292 | -34.8221 | -24.7477
15.5353 | -34.8226 | -24.7559
15.5414 | -34.8231 | -24.7639
15.5476 | -34.8236 | -24.7717
15.5537 | -34.824 | -24.7797
15.5598 | -34.8245 | -24.7879
15.5659 | -34.825 | -24.7959
15.572 | -34.8255 | -24.8039
15.5781 | -34.826 | -24.8117
15.5842 | -34.8265 | -24.8196
15.5903 | -34.827 | -24.8274
15.5964 | -34.8275 | -24.8353
15.6025 | -34.828 | -24.8432
15.6086 | -34.8285 | -24.8512
15.6147 | -34.8289 | -24.8594
15.6208 | -34.8294 | -24.8671
15.6269 | -34.8299 | -24.875
15.633 | -34.8304 | -24.883
15.6391 | -34.8309 | -24.8911
15.6452 | -34.8314 | -24.8991
15.6513 | -34.8319 | -24.9067
15.6574 | -34.8324 | -24.9148
15.6635 | -34.8329 | -24.9227
15.6696 | -34.8334 | -24.9306
15.6757 | -34.8339 | -24.9386
15.6818 | -34.8343 | -24.9465
15.688 | -34.8348 | -24.9544
15.6941 | -34.8353 | -24.9623
15.7002 | -34.8358 | -24.9702
15.7063 | -34.8363 | -24.9779
15.7124 | -34.8368 | -24.986
15.7185 | -34.8372 | -24.994
15.7246 | -34.8378 | -25.0019
15.7307 | -34.8383 | -25.0098
15.7368 | -34.8388 | -25.0179
15.7429 | -34.8393 | -25.0257
15.749 | -34.8398 | -25.0337
15.7551 | -34.8403 | -25.0418
15.7612 | -34.8408 | -25.0496
15.7673 | -34.8413 | -25.0576
15.7734 | -34.8418 | -25.0655
15.7795 | -34.8423 | -25.0734
15.7856 | -34.8427 | -25.0812
15.7917 | -34.8433 | -25.0889
15.7978 | -34.8437 | -25.097
15.8039 | -34.8442 | -25.1048
15.81 | -34.8447 | -25.1127
15.8161 | -34.8452 | -25.1206
15.8222 | -34.8457 | -25.1285
15.8283 | -34.8463 | -25.1361
15.8345 | -34.8467 | -25.144
15.8406 | -34.8472 | -25.1518
15.8467 | -34.8477 | -25.1597
15.8528 | -34.8482 | -25.1676
15.8589 | -34.8487 | -25.1755
15.865 | -34.8493 | -25.1831
15.8711 | -34.8498 | -25.1911
15.8772 | -34.8503 | -25.1988
15.8833 | -34.8507 | -25.2068
15.8894 | -34.8512 | -25.2146
15.8955 | -34.8517 | -25.2225
15.9016 | -34.8522 | -25.23
15.9077 | -34.8527 | -25.2381
15.9138 | -34.8531 | -25.246
15.9199 | -34.8537 | -25.2538
15.926 | -34.8541 | -25.2615
15.9321 | -34.8547 | -25.2691
15.9382 | -34.8551 | -25.2769
15.9443 | -34.8556 | -25.2847
15.9504 | -34.8561 | -25.2927
15.9565 | -34.8566 | -25.3005
15.9626 | -34.857 | -25.3084
15.9687 | -34.8575 | -25.3159
15.9749 | -34.858 | -25.3235
15.981 | -34.8585 | -25.3313
15.9871 | -34.859 | -25.3392
15.9932 | -34.8595 | -25.3469
15.9993 | -34.86 | -25.3547
16.0054 | -34.8605 | -25.3623
16.0115 | -34.861 | -25.3702
16.0176 | -34.8615 | -25.3781
16.0237 | -34.862 | -25.3861
16.0298 | -34.8625 | -25.3938
16.0359 | -34.863 | -25.4017
16.042 | -34.8635 | -25.4093
16.0481 | -34.864 | -25.4172
16.0542 | -34.8645 | -25.4251
16.0603 | -34.865 | -25.4328
16.0664 | -34.8655 | -25.4407
16.0725 | -34.866 | -25.4483
16.0786 | -34.8665 | -25.4561
16.0847 | -34.867 | -25.4639
16.0908 | -34.8674 | -25.4717
16.0969 | -34.8679 | -25.4795
16.103 | -34.8684 | -25.4874
16.1091 | -34.8689 | -25.495
16.1152 | -34.8694 | -25.5028
16.1214 | -34.8699 | -25.5107
16.1275 | -34.8704 | -25.5184
16.1336 | -34.8709 | -25.5261
16.1397 | -34.8714 | -25.5339
16.1458 | -34.8719 | -25.5415
16.1519 | -34.8724 | -25.5494
16.158 | -34.8729 | -25.5573
16.1641 | -34.8734 | -25.5653
16.1702 | -34.8739 | -25.5731
16.1763 | -34.8744 | -25.5809
16.1824 | -34.8749 | -25.5884
16.1885 | -34.8754 | -25.596
16.1946 | -34.8759 | -25.6037
16.2007 | -34.8764 | -25.6114
16.2068 | -34.8769 | -25.6191
16.2129 | -34.8774 | -25.627
16.219 | -34.8779 | -25.6344
16.2251 | -34.8784 | -25.642
16.2312 | -34.8789 | -25.6498
16.2373 | -34.8794 | -25.6575
16.2434 | -34.8799 | -25.6652
16.2495 | -34.8805 | -25.6726
16.2556 | -34.881 | -25.6806
16.2618 | -34.8815 | -25.6883
16.2679 | -34.882 | -25.696
16.274 | -34.8825 | -25.7038
16.2801 | -34.883 | -25.7114
16.2862 | -34.8835 | -25.719
16.2923 | -34.884 | -25.7266
16.2984 | -34.8844 | -25.7344
16.3045 | -34.8849 | -25.7422
16.3106 | -34.8854 | -25.75
16.3167 | -34.8859 | -25.7577
16.3228 | -34.8864 | -25.7651
16.3289 | -34.8869 | -25.7728
16.335 | -34.8874 | -25.7803
16.3411 | -34.8879 | -25.788
16.3472 | -34.8884 | -25.7959
16.3533 | -34.8889 | -25.8036
16.3594 | -34.8894 | -25.811
16.3655 | -34.8899 | -25.8186
16.3716 | -34.8903 | -25.8261
16.3777 | -34.8909 | -25.8337
16.3838 | -34.8913 | -25.8412
16.3899 | -34.8918 | -25.8488
16.396 | -34.8923 | -25.8565
16.4021 | -34.8928 | -25.864
16.4083 | -34.8933 | -25.8716
16.4144 | -34.8938 | -25.8794
16.4205 | -34.8943 | -25.887
16.4266 | -34.8948 | -25.8943
16.4327 | -34.8953 | -25.9019
16.4388 | -34.8958 | -25.9096
16.4449 | -34.8963 | -25.9172
16.451 | -34.8968 | -25.9249
16.4571 | -34.8973 | -25.9327
16.4632 | -34.8977 | -25.9398
16.4693 | -34.8982 | -25.9475
16.4754 | -34.8987 | -25.9553
16.4815 | -34.8992 | -25.9629
16.4876 | -34.8997 | -25.9703
16.4937 | -34.9002 | -25.9778
16.4998 | -34.9007 | -25.9851
16.5059 | -34.9012 | -25.9926
16.512 | -34.9017 | -26.0004
16.5181 | -34.9022 | -26.0081
16.5242 | -34.9027 | -26.0159
16.5303 | -34.9032 | -26.0233
16.5364 | -34.9037 | -26.0308
16.5425 | -34.9042 | -26.0385
16.5487 | -34.9047 | -26.046
16.5548 | -34.9051 | -26.0537
16.5609 | -34.9056 | -26.0612
16.567 | -34.9061 | -26.0686
16.5731 | -34.9066 | -26.0764
16.5792 | -34.9071 | -26.0842
16.5853 | -34.9076 | -26.0919
16.5914 | -34.9081 | -26.0996
16.5975 | -34.9086 | -26.1072
16.6036 | -34.9091 | -26.1146
16.6097 | -34.9095 | -26.1223
16.6158 | -34.91 | -26.1297
16.6219 | -34.9105 | -26.1374
16.628 | -34.911 | -26.145
16.6341 | -34.9115 | -26.1525
16.6402 | -34.9119 | -26.16
16.6463 | -34.9124 | -26.1676
16.6524 | -34.9129 | -26.1751
16.6585 | -34.9134 | -26.1826
16.6646 | -34.9139 | -26.1901
16.6707 | -34.9144 | -26.1976
16.6768 | -34.9149 | -26.2051
16.6829 | -34.9154 | -26.2125
16.689 | -34.9159 | -26.2202
16.6952 | -34.9163 | -26.2277
16.7013 | -34.9168 | -26.2352
16.7074 | -34.9173 | -26.2425
16.7135 | -34.9178 | -26.2501
16.7196 | -34.9183 | -26.2574
16.7257 | -34.9188 | -26.2651
16.7318 | -34.9193 | -26.2727
16.7379 | -34.9198 | -26.2802
16.744 | -34.9203 | -26.2874
16.7501 | -34.9208 | -26.2949
16.7562 | -34.9213 | -26.3025
16.7623 | -34.9218 | -26.3099
16.7684 | -34.9222 | -26.3174
16.7745 | -34.9227 | -26.325
16.7806 | -34.9232 | -26.3323
16.7867 | -34.9237 | -26.3398
16.7928 | -34.9242 | -26.3474
16.7989 | -34.9247 | -26.3548
16.805 | -34.9252 | -26.3624
16.8111 | -34.9257 | -26.3698
16.8172 | -34.9262 | -26.3774
16.8233 | -34.9267 | -26.385
16.8294 | -34.9272 | -26.3926
16.8356 | -34.9276 | -26.4002
16.8417 | -34.9282 | -26.4077
16.8478 | -34.9287 | -26.4152
16.8539 | -34.9291 | -26.4225
16.86 | -34.9296 | -26.4301
16.8661 | -34.9301 | -26.4375
16.8722 | -34.9306 | -26.4452
16.8783 | -34.9311 | -26.4528
16.8844 | -34.9316 | -26.4603
16.8905 | -34.9321 | -26.4678
16.8966 | -34.9325 | -26.4751
16.9027 | -34.933 | -26.4827
16.9088 | -34.9335 | -26.4903
16.9149 | -34.934 | -26.4979
16.921 | -34.9345 | -26.5052
16.9271 | -34.935 | -26.5127
16.9332 | -34.9355 | -26.5202
16.9393 | -34.936 | -26.5279
16.9454 | -34.9365 | -26.5352
16.9515 | -34.9369 | -26.5427
16.9576 | -34.9374 | -26.5499
16.9637 | -34.9379 | -26.5573
16.9698 | -34.9384 | -26.5647
16.9759 | -34.9389 | -26.5723
16.9821 | -34.9393 | -26.58
16.9882 | -34.9398 | -26.5873
16.9943 | -34.9403 | -26.5947
17.0004 | -34.9408 | -26.6021
17.0065 | -34.9413 | -26.6097
17.0126 | -34.9418 | -26.6172
17.0187 | -34.9423 | -26.6249
17.0248 | -34.9428 | -26.6324
17.0309 | -34.9432 | -26.6398
17.037 | -34.9437 | -26.6473
17.0431 | -34.9442 | -26.6546
17.0492 | -34.9447 | -26.6621
17.0553 | -34.9452 | -26.6696
17.0614 | -34.9457 | -26.677
17.0675 | -34.9462 | -26.6844
17.0736 | -34.9466 | -26.6918
17.0797 | -34.9471 | -26.6993
17.0858 | -34.9476 | -26.7069
17.0919 | -34.9481 | -26.7143
17.098 | -34.9486 | -26.7217
17.1041 | -34.949 | -26.7292
17.1102 | -34.9496 | -26.7367
17.1163 | -34.95 | -26.744
17.1225 | -34.9505 | -26.7514
17.1286 | -34.951 | -26.7585
17.1347 | -34.9515 | -26.7661
17.1408 | -34.952 | -26.7736
17.1469 | -34.9524 | -26.7811
17.153 | -34.9529 | -26.7885
17.1591 | -34.9534 | -26.7958
17.1652 | -34.9539 | -26.8032
17.1713 | -34.9543 | -26.8105
17.1774 | -34.9548 | -26.818
17.1835 | -34.9553 | -26.8252
17.1896 | -34.9558 | -26.8327
17.1957 | -34.9563 | -26.8401
17.2018 | -34.9568 | -26.8471
17.2079 | -34.9573 | -26.8546
17.214 | -34.9577 | -26.8622
17.2201 | -34.9582 | -26.8695
17.2262 | -34.9587 | -26.8769
17.2323 | -34.9592 | -26.8844
17.2384 | -34.9597 | -26.8916
17.2445 | -34.9602 | -26.899
17.2506 | -34.9606 | -26.9063
17.2567 | -34.9611 | -26.9137
17.2628 | -34.9616 | -26.9212
17.269 | -34.9621 | -26.9286
17.2751 | -34.9625 | -26.9361
17.2812 | -34.963 | -26.9436
17.2873 | -34.9635 | -26.9509
17.2934 | -34.964 | -26.9583
17.2995 | -34.9645 | -26.9658
17.3056 | -34.965 | -26.9731
17.3117 | -34.9655 | -26.9804
17.3178 | -34.9659 | -26.9878
17.3239 | -34.9664 | -26.9952
17.33 | -34.9669 | -27.0027
17.3361 | -34.9674 | -27.0101
17.3422 | -34.9678 | -27.0173
17.3483 | -34.9683 | -27.0247
17.3544 | -34.9688 | -27.0321
17.3605 | -34.9693 | -27.0396
17.3666 | -34.9697 | -27.047
17.3727 | -34.9702 | -27.0544
17.3788 | -34.9707 | -27.0616
17.3849 | -34.9712 | -27.0689
17.391 | -34.9717 | -27.0762
17.3971 | -34.9722 | -27.0836
17.4032 | -34.9726 | -27.0909
17.4094 | -34.9731 | -27.0982
17.4155 | -34.9736 | -27.1055
17.4216 | -34.9741 | -27.1128
17.4277 | -34.9746 | -27.1203
17.4338 | -34.975 | -27.1276
17.4399 | -34.9755 | -27.135
17.446 | -34.976 | -27.1421
17.4521 | -34.9765 | -27.1495
17.4582 | -34.977 | -27.1567
17.4643 | -34.9774 | -27.1641
17.4704 | -34.9779 | -27.1714
17.4765 | -34.9784 | -27.1789
17.4826 | -34.9788 | -27.1861
17.4887 | -34.9793 | -27.1935
17.4948 | -34.9798 | -27.2008
17.5009 | -34.9803 | -27.2081
17.507 | -34.9807 | -27.2157
17.5131 | -34.9813 | -27.223
17.5192 | -34.9817 | -27.2299
17.5253 | -34.9822 | -27.2371
17.5314 | -34.9826 | -27.2445
17.5375 | -34.9831 | -27.2518
17.5436 | -34.9836 | -27.2593
17.5497 | -34.9841 | -27.2667
17.5559 | -34.9846 | -27.2738
17.562 | -34.9851 | -27.2809
17.5681 | -34.9856 | -27.2882
17.5742 | -34.986 | -27.2957
17.5803 | -34.9865 | -27.3031
17.5864 | -34.987 | -27.3104
17.5925 | -34.9875 | -27.3178
17.5986 | -34.988 | -27.325
17.6047 | -34.9885 | -27.3324
17.6108 | -34.9889 | -27.3396
17.6169 | -34.9894 | -27.347
17.623 | -34.9899 | -27.3541
17.6291 | -34.9904 | -27.3614
17.6352 | -34.9909 | -27.3689
17.6413 | -34.9913 | -27.3763
17.6474 | -34.9918 | -27.3837
17.6535 | -34.9923 | -27.391
17.6596 | -34.9928 | -27.3981
17.6657 | -34.9933 | -27.4055
17.6718 | -34.9938 | -27.4129
17.6779 | -34.9942 | -27.4201
17.684 | -34.9947 | -27.4275
17.6901 | -34.9952 | -27.4349
17.6963 | -34.9956 | -27.4421
17.7024 | -34.9961 | -27.4493
17.7085 | -34.9966 | -27.4564
17.7146 | -34.9971 | -27.4636
17.7207 | -34.9975 | -27.4708
17.7268 | -34.998 | -27.478
17.7329 | -34.9985 | -27.4855
17.739 | -34.9989 | -27.4929
17.7451 | -34.9994 | -27.5
17.7512 | -34.9999 | -27.5072
17.7573 | -35.0004 | -27.5147
17.7634 | -35.0009 | -27.5217
17.7695 | -35.0014 | -27.529
17.7756 | -35.0019 | -27.5364
17.7817 | -35.0023 | -27.5437
17.7878 | -35.0028 | -27.5512
17.7939 | -35.0033 | -27.5585
17.8 | -35.0037 | -27.5658
17.8061 | -35.0042 | -27.5731
17.8122 | -35.0047 | -27.5805
17.8183 | -35.0052 | -27.5878
17.8244 | -35.0056 | -27.5951
17.8305 | -35.0061 | -27.6024
17.8366 | -35.0066 | -27.6095
17.8428 | -35.007 | -27.6168
17.8489 | -35.0075 | -27.6241
17.855 | -35.008 | -27.6313
17.8611 | -35.0085 | -27.6385
17.8672 | -35.009 | -27.6456
17.8733 | -35.0094 | -27.653
17.8794 | -35.0099 | -27.6602
17.8855 | -35.0104 | -27.6673
17.8916 | -35.0108 | -27.6746
17.8977 | -35.0113 | -27.6819
17.9038 | -35.0118 | -27.6891
17.9099 | -35.0122 | -27.6964
17.916 | -35.0127 | -27.7038
17.9221 | -35.0132 | -27.7111
17.9282 | -35.0137 | -27.7182
17.9343 | -35.0141 | -27.7256
17.9404 | -35.0146 | -27.7327
17.9465 | -35.0151 | -27.7399
17.9526 | -35.0155 | -27.7472
17.9587 | -35.016 | -27.7544
17.9648 | -35.0165 | -27.7616
17.9709 | -35.017 | -27.769
17.977 | -35.0174 | -27.7761
17.9832 | -35.0179 | -27.7835
17.9893 | -35.0184 | -27.7909
17.9954 | -35.0188 | -27.7981
18.0015 | -35.0193 | -27.8055
18.0076 | -35.0197 | -27.8124
18.0137 | -35.0202 | -27.8197
18.0198 | -35.0206 | -27.8268
18.0259 | -35.0211 | -27.8339
18.032 | -35.0216 | -27.8413
18.0381 | -35.0221 | -27.8488
18.0442 | -35.0225 | -27.856
18.0503 | -35.023 | -27.8634
18.0564 | -35.0234 | -27.8707
18.0625 | -35.0239 | -27.8781
18.0686 | -35.0244 | -27.8853
18.0747 | -35.0248 | -27.8925
18.0808 | -35.0253 | -27.8996
18.0869 | -35.0258 | -27.907
18.093 | -35.0263 | -27.9143
18.0991 | -35.0267 | -27.9215
18.1052 | -35.0272 | -27.9287
18.1113 | -35.0277 | -27.936
18.1174 | -35.0281 | -27.9431
18.1236 | -35.0286 | -27.9504
18.1297 | -35.0291 | -27.9576
18.1358 | -35.0295 | -27.9648
18.1419 | -35.03 | -27.9719
18.148 | -35.0305 | -27.9791
18.1541 | -35.031 | -27.9861
18.1602 | -35.0314 | -27.9934
18.1663 | -35.0319 | -28.0006
18.1724 | -35.0324 | -28.0079
18.1785 | -35.0328 | -28.0151
18.1846 | -35.0333 | -28.0224
18.1907 | -35.0337 | -28.0294
18.1968 | -35.0342 | -28.0366
18.2029 | -35.0347 | -28.044
18.209 | -35.0351 | -28.0514
18.2151 | -35.0356 | -28.0588
18.2212 | -35.0361 | -28.066
18.2273 | -35.0366 | -28.0733
18.2334 | -35.0371 | -28.0806
18.2395 | -35.0375 | -28.0879
18.2456 | -35.038 | -28.0953
18.2517 | -35.0384 | -28.1025
18.2578 | -35.0389 | -28.1095
18.2639 | -35.0394 | -28.1169
18.2701 | -35.0398 | -28.124
18.2762 | -35.0403 | -28.1311
18.2823 | -35.0408 | -28.1381
18.2884 | -35.0412 | -28.1452
18.2945 | -35.0417 | -28.1524
18.3006 | -35.0422 | -28.1597
18.3067 | -35.0426 | -28.1668
18.3128 | -35.0431 | -28.1741
18.3189 | -35.0435 | -28.1814
18.325 | -35.044 | -28.1884
18.3311 | -35.0444 | -28.1956
18.3372 | -35.0449 | -28.2027
18.3433 | -35.0453 | -28.21
18.3494 | -35.0458 | -28.2173
18.3555 | -35.0463 | -28.2244
18.3616 | -35.0468 | -28.2315
18.3677 | -35.0473 | -28.2387
18.3738 | -35.0478 | -28.2458
18.3799 | -35.0483 | -28.253
18.386 | -35.0487 | -28.2603
18.3921 | -35.0491 | -28.2673
18.3982 | -35.0496 | -28.2745
18.4043 | -35.05 | -28.2817
18.4105 | -35.0505 | -28.2889
18.4166 | -35.0509 | -28.2961
18.4227 | -35.0514 | -28.3031
18.4288 | -35.0518 | -28.3104
18.4349 | -35.0523 | -28.3176
18.441 | -35.0527 | -28.3247
18.4471 | -35.0531 | -28.3319
18.4532 | -35.0536 | -28.3393
18.4593 | -35.054 | -28.3466
18.4654 | -35.0545 | -28.3537
18.4715 | -35.0549 | -28.3609
18.4776 | -35.0554 | -28.3682
18.4837 | -35.0558 | -28.3755
18.4898 | -35.0563 | -28.3829
18.4959 | -35.0568 | -28.39
18.502 | -35.0572 | -28.3975
18.5081 | -35.0576 | -28.4046
18.5142 | -35.0581 | -28.4118
18.5203 | -35.0585 | -28.4188
18.5264 | -35.0589 | -28.4261
18.5325 | -35.0594 | -28.4334
18.5386 | -35.0599 | -28.4403
18.5447 | -35.0604 | -28.4475
18.5508 | -35.0608 | -28.4544
18.557 | -35.0613 | -28.4616
18.5631 | -35.0617 | -28.4686
18.5692 | -35.0622 | -28.4758
18.5753 | -35.0626 | -28.4831
18.5814 | -35.0631 | -28.4904
18.5875 | -35.0635 | -28.4976
18.5936 | -35.0639 | -28.5046
18.5997 | -35.0644 | -28.512
18.6058 | -35.0648 | -28.5191
18.6119 | -35.0653 | -28.5265
18.618 | -35.0657 | -28.5335
18.6241 | -35.0662 | -28.5408
18.6302 | -35.0667 | -28.5481
18.6363 | -35.0672 | -28.5551
18.6424 | -35.0676 | -28.5622
18.6485 | -35.0681 | -28.5694
18.6546 | -35.0685 | -28.5763
18.6607 | -35.069 | -28.5838
18.6668 | -35.0694 | -28.591
18.6729 | -35.0699 | -28.5981
18.679 | -35.0704 | -28.6052
18.6851 | -35.0709 | -28.612
18.6912 | -35.0713 | -28.6189
18.6974 | -35.0718 | -28.6258
18.7035 | -35.0722 | -28.633
18.7096 | -35.0727 | -28.6402
18.7157 | -35.0731 | -28.6476
18.7218 | -35.0736 | -28.6548
18.7279 | -35.0741 | -28.6621
18.734 | -35.0746 | -28.6694
18.7401 | -35.0751 | -28.6766
18.7462 | -35.0756 | -28.6836
18.7523 | -35.0761 | -28.6907
18.7584 | -35.0765 | -28.6978
18.7645 | -35.077 | -28.7047
18.7706 | -35.0775 | -28.7121
18.7767 | -35.0779 | -28.7192
18.7828 | -35.0784 | -28.7263
18.7889 | -35.0788 | -28.7333
18.795 | -35.0793 | -28.7407
18.8011 | -35.0797 | -28.7478
18.8072 | -35.0802 | -28.7551
18.8133 | -35.0807 | -28.7623
18.8194 | -35.0811 | -28.7697
18.8255 | -35.0815 | -28.7767
18.8316 | -35.082 | -28.784
18.8377 | -35.0824 | -28.7912
18.8439 | -35.0829 | -28.7983
18.85 | -35.0833 | -28.8058
18.8561 | -35.0838 | -28.8129
18.8622 | -35.0843 | -28.8201
18.8683 | -35.0847 | -28.827
18.8744 | -35.0852 | -28.8336
18.8805 | -35.0857 | -28.8409
18.8866 | -35.0862 | -28.8483
18.8927 | -35.0867 | -28.8551
18.8988 | -35.0872 | -28.8621
18.9049 | -35.0876 | -28.869
18.911 | -35.088 | -28.8763
18.9171 | -35.0885 | -28.8834
18.9232 | -35.089 | -28.8904
18.9293 | -35.0895 | -28.8974
18.9354 | -35.09 | -28.9044
18.9415 | -35.0905 | -28.9114
18.9476 | -35.0909 | -28.9184
18.9537 | -35.0914 | -28.9257
18.9598 | -35.0918 | -28.9321
18.9659 | -35.0923 | -28.9391
18.972 | -35.0927 | -28.946
18.9781 | -35.0932 | -28.9532
18.9843 | -35.0936 | -28.9604
18.9904 | -35.0941 | -28.9674
18.9965 | -35.0946 | -28.9743
19.0026 | -35.0951 | -28.9815
19.0087 | -35.0955 | -28.9887
19.0148 | -35.096 | -28.9956
19.0209 | -35.0964 | -29.0028
19.027 | -35.0969 | -29.0102
19.0331 | -35.0974 | -29.017
19.0392 | -35.0978 | -29.0242
19.0453 | -35.0983 | -29.031
19.0514 | -35.0988 | -29.0379
19.0575 | -35.0992 | -29.0451
19.0636 | -35.0996 | -29.0517
19.0697 | -35.1001 | -29.0586
19.0758 | -35.1005 | -29.0656
19.0819 | -35.101 | -29.0727
19.088 | -35.1015 | -29.0797
19.0941 | -35.1019 | -29.0869
19.1002 | -35.1024 | -29.0938
19.1063 | -35.1028 | -29.1008
19.1124 | -35.1032 | -29.1079
19.1185 | -35.1037 | -29.1149
19.1246 | -35.1041 | -29.1221
19.1308 | -35.1046 | -29.1294
19.1369 | -35.105 | -29.1361
19.143 | -35.1055 | -29.1431
19.1491 | -35.1059 | -29.1503
19.1552 | -35.1063 | -29.1574
19.1613 | -35.1067 | -29.1644
19.1674 | -35.1072 | -29.1716
19.1735 | -35.1076 | -29.1788
19.1796 | -35.1081 | -29.1856
19.1857 | -35.1085 | -29.1928
19.1918 | -35.109 | -29.1996
19.1979 | -35.1094 | -29.207
19.204 | -35.1099 | -29.214
19.2101 | -35.1103 | -29.221
19.2162 | -35.1107 | -29.2281
19.2223 | -35.1112 | -29.235
19.2284 | -35.1117 | -29.242
19.2345 | -35.1121 | -29.2489
19.2406 | -35.1126 | -29.2557
19.2467 | -35.113 | -29.2629
19.2528 | -35.1134 | -29.27
19.2589 | -35.1139 | -29.2771
19.265 | -35.1144 | -29.2843
19.2712 | -35.1148 | -29.2912
19.2773 | -35.1153 | -29.2981
19.2834 | -35.1158 | -29.3053
19.2895 | -35.1163 | -29.3124
19.2956 | -35.1167 | -29.3195
19.3017 | -35.1172 | -29.3265
19.3078 | -35.1176 | -29.3339
19.3139 | -35.1181 | -29.3408
19.32 | -35.1185 | -29.3478
19.3261 | -35.1189 | -29.3549
19.3322 | -35.1194 | -29.3616
19.3383 | -35.1199 | -29.3684
19.3444 | -35.1203 | -29.3753
19.3505 | -35.1208 | -29.3824
19.3566 | -35.1212 | -29.3893
19.3627 | -35.1217 | -29.3964
19.3688 | -35.1222 | -29.4032
19.3749 | -35.1226 | -29.4103
19.381 | -35.1231 | -29.4172
19.3871 | -35.1235 | -29.424
19.3932 | -35.1239 | -29.4311
19.3993 | -35.1243 | -29.438
19.4054 | -35.1247 | -29.4449
19.4115 | -35.1252 | -29.4519
19.4177 | -35.1256 | -29.459
19.4238 | -35.1261 | -29.4659
19.4299 | -35.1265 | -29.473
19.436 | -35.127 | -29.4801
19.4421 | -35.1274 | -29.4871
19.4482 | -35.1278 | -29.494
19.4543 | -35.1283 | -29.5008
19.4604 | -35.1287 | -29.5079
19.4665 | -35.1292 | -29.5148
19.4726 | -35.1296 | -29.5217
19.4787 | -35.1301 | -29.5285
19.4848 | -35.1305 | -29.5357
19.4909 | -35.131 | -29.5429
19.497 | -35.1314 | -29.55
19.5031 | -35.1318 | -29.5568
19.5092 | -35.1322 | -29.5639
19.5153 | -35.1327 | -29.571
19.5214 | -35.1332 | -29.5783
19.5275 | -35.1336 | -29.5851
19.5336 | -35.134 | -29.5921
19.5397 | -35.1345 | -29.599
19.5458 | -35.135 | -29.6058
19.5519 | -35.1355 | -29.613
19.5581 | -35.1358 | -29.6196
19.5642 | -35.1363 | -29.6266
19.5703 | -35.1367 | -29.6334
19.5764 | -35.1372 | -29.6404
19.5825 | -35.1376 | -29.6472
19.5886 | -35.138 | -29.654
19.5947 | -35.1385 | -29.6609
19.6008 | -35.1389 | -29.6677
19.6069 | -35.1393 | -29.6745
19.613 | -35.1398 | -29.6817
19.6191 | -35.1403 | -29.6886
19.6252 | -35.1407 | -29.6956
19.6313 | -35.1411 | -29.7025
19.6374 | -35.1415 | -29.7097
19.6435 | -35.1419 | -29.7165
19.6496 | -35.1424 | -29.7234
19.6557 | -35.1428 | -29.7304
19.6618 | -35.1432 | -29.737
19.6679 | -35.1437 | -29.7439
19.674 | -35.1441 | -29.7508
19.6801 | -35.1445 | -29.7575
19.6862 | -35.145 | -29.7646
19.6923 | -35.1455 | -29.7715
19.6984 | -35.1459 | -29.7778
19.7046 | -35.1464 | -29.7846
19.7107 | -35.1468 | -29.7913
19.7168 | -35.1472 | -29.7983
19.7229 | -35.1477 | -29.8053
19.729 | -35.1481 | -29.8122
19.7351 | -35.1486 | -29.8188
19.7412 | -35.149 | -29.8262
19.7473 | -35.1495 | -29.8333
19.7534 | -35.1499 | -29.8405
19.7595 | -35.1503 | -29.8474
19.7656 | -35.1508 | -29.8545
19.7717 | -35.1513 | -29.8609
19.7778 | -35.1517 | -29.8678
19.7839 | -35.1522 | -29.8749
19.79 | -35.1527 | -29.8818
19.7961 | -35.1531 | -29.8888
19.8022 | -35.1535 | -29.8954
19.8083 | -35.154 | -29.9023
19.8144 | -35.1545 | -29.9095
19.8205 | -35.155 | -29.9166
19.8266 | -35.1554 | -29.9235
19.8327 | -35.1558 | -29.9305
19.8388 | -35.1563 | -29.9369
19.845 | -35.1567 | -29.9435
19.8511 | -35.1571 | -29.9503
19.8572 | -35.1576 | -29.957
19.8633 | -35.158 | -29.9639
19.8694 | -35.1584 | -29.9708
19.8755 | -35.1589 | -29.9773
19.8816 | -35.1593 | -29.9845
19.8877 | -35.1598 | -29.9918
19.8938 | -35.1603 | -29.9988
19.8999 | -35.1608 | -30.0059
19.906 | -35.1611 | -30.0125
19.9121 | -35.1616 | -30.0191
19.9182 | -35.1621 | -30.0261
19.9243 | -35.1625 | -30.0331
19.9304 | -35.1629 | -30.04
19.9365 | -35.1633 | -30.047
19.9426 | -35.1638 | -30.0536
19.9487 | -35.1642 | -30.0604
19.9548 | -35.1646 | -30.0676
19.9609 | -35.1651 | -30.074
19.967 | -35.1656 | -30.0806
19.9731 | -35.166 | -30.0873
19.9792 | -35.1665 | -30.0942
19.9853 | -35.167 | -30.1011
19.9915 | -35.1675 | -30.1079
19.9976 | -35.1679 | -30.1146
20.0037 | -35.1683 | -30.1214
20.0098 | -35.1687 | -30.1282
20.0159 | -35.1691 | -30.1349
20.022 | -35.1695 | -30.1417
20.0281 | -35.1699 | -30.1485
20.0342 | -35.1703 | -30.1554
20.0403 | -35.1708 | -30.1624
20.0464 | -35.1712 | -30.1694
20.0525 | -35.1717 | -30.1758
20.0586 | -35.1721 | -30.1828
20.0647 | -35.1725 | -30.1895
20.0708 | -35.173 | -30.1962
20.0769 | -35.1735 | -30.2029
20.083 | -35.1739 | -30.2092
20.0891 | -35.1743 | -30.2159
20.0952 | -35.1747 | -30.2227
20.1013 | -35.1752 | -30.2293
20.1074 | -35.1756 | -30.2361
20.1135 | -35.1761 | -30.2431
20.1196 | -35.1765 | -30.2497
20.1257 | -35.1769 | -30.2566
20.1319 | -35.1774 | -30.2633
20.138 | -35.1778 | -30.2701
20.1441 | -35.1783 | -30.2767
20.1502 | -35.1787 | -30.2837
20.1563 | -35.1791 | -30.2899
20.1624 | -35.1795 | -30.2966
20.1685 | -35.18 | -30.3038
20.1746 | -35.1804 | -30.3105
20.1807 | -35.1809 | -30.3174
20.1868 | -35.1814 | -30.3243
20.1929 | -35.1818 | -30.331
20.199 | -35.1822 | -30.338
20.2051 | -35.1827 | -30.3449
20.2112 | -35.1831 | -30.3517
20.2173 | -35.1836 | -30.3585
20.2234 | -35.184 | -30.3653
20.2295 | -35.1844 | -30.3721
20.2356 | -35.1848 | -30.3794
20.2417 | -35.1852 | -30.3858
20.2478 | -35.1857 | -30.392
20.2539 | -35.1861 | -30.3986
20.26 | -35.1866 | -30.4051
20.2661 | -35.1871 | -30.4118
20.2723 | -35.1875 | -30.4186
20.2784 | -35.1879 | -30.4254
20.2845 | -35.1884 | -30.4323
20.2906 | -35.1889 | -30.4391
20.2967 | -35.1893 | -30.4456
20.3028 | -35.1897 | -30.4525
20.3089 | -35.1902 | -30.4593
20.315 | -35.1907 | -30.4658
20.3211 | -35.1911 | -30.4726
20.3272 | -35.1915 | -30.4794
20.3333 | -35.192 | -30.4857
20.3394 | -35.1924 | -30.4922
20.3455 | -35.1929 | -30.4988
20.3516 | -35.1933 | -30.505
20.3577 | -35.1937 | -30.512
20.3638 | -35.1942 | -30.5186
20.3699 | -35.1947 | -30.5245
20.376 | -35.1951 | -30.5313
20.3821 | -35.1956 | -30.538
20.3882 | -35.196 | -30.5448
20.3943 | -35.1965 | -30.5518
20.4004 | -35.1969 | -30.5585
20.4065 | -35.1973 | -30.5649
20.4126 | -35.1977 | -30.5719
20.4188 | -35.1982 | -30.5783
20.4249 | -35.1986 | -30.585
20.431 | -35.1991 | -30.592
20.4371 | -35.1997 | -30.5981
20.4432 | -35.2001 | -30.605
20.4493 | -35.2005 | -30.612
20.4554 | -35.2009 | -30.6185
20.4615 | -35.2014 | -30.6253
20.4676 | -35.2018 | -30.6316
20.4737 | -35.2023 | -30.6382
20.4798 | -35.2028 | -30.6446
20.4859 | -35.2032 | -30.6514
20.492 | -35.2036 | -30.6579
20.4981 | -35.204 | -30.6645
20.5042 | -35.2045 | -30.6712
20.5103 | -35.2049 | -30.6772
20.5164 | -35.2053 | -30.6837
20.5225 | -35.2057 | -30.6908
20.5286 | -35.2062 | -30.6976
20.5347 | -35.2066 | -30.7043
20.5408 | -35.207 | -30.711
20.5469 | -35.2075 | -30.7178
20.553 | -35.2079 | -30.7245
20.5592 | -35.2084 | -30.7307
20.5653 | -35.2088 | -30.7372
20.5714 | -35.2092 | -30.7437
20.5775 | -35.2097 | -30.75
20.5836 | -35.2101 | -30.7565
20.5897 | -35.2106 | -30.7629
20.5958 | -35.2111 | -30.7697
20.6019 | -35.2115 | -30.7763
20.608 | -35.212 | -30.7831
20.6141 | -35.2124 | -30.7896
20.6202 | -35.2129 | -30.7959
20.6263 | -35.2133 | -30.8027
20.6324 | -35.2138 | -30.8098
20.6385 | -35.2142 | -30.8162
20.6446 | -35.2147 | -30.8227
20.6507 | -35.2151 | -30.8292
20.6568 | -35.2155 | -30.8355
20.6629 | -35.216 | -30.8422
20.669 | -35.2164 | -30.8486
20.6751 | -35.2168 | -30.8552
20.6812 | -35.2172 | -30.8614
20.6873 | -35.2177 | -30.8682
20.6934 | -35.2182 | -30.8749
20.6995 | -35.2186 | -30.8815
20.7057 | -35.219 | -30.8883
20.7118 | -35.2195 | -30.8949
20.7179 | -35.22 | -30.9009
20.724 | -35.2204 | -30.9076
20.7301 | -35.2208 | -30.9143
20.7362 | -35.2212 | -30.9213
20.7423 | -35.2217 | -30.9274
20.7484 | -35.2221 | -30.9341
20.7545 | -35.2226 | -30.9397
20.7606 | -35.2231 | -30.9465
20.7667 | -35.2235 | -30.9533
20.7728 | -35.224 | -30.9595
20.7789 | -35.2245 | -30.9663
20.785 | -35.2249 | -30.9727
20.7911 | -35.2254 | -30.9787
20.7972 | -35.2258 | -30.9845
20.8033 | -35.2263 | -30.9908
20.8094 | -35.2268 | -30.9974
20.8155 | -35.2273 | -31.0041
20.8216 | -35.2277 | -31.0102
20.8277 | -35.2282 | -31.016
20.8338 | -35.2287 | -31.0224
20.8399 | -35.2292 | -31.0289
20.8461 | -35.2296 | -31.0354
20.8522 | -35.2301 | -31.0419
20.8583 | -35.2305 | -31.0483
20.8644 | -35.231 | -31.0549
20.8705 | -35.2314 | -31.0618
20.8766 | -35.2318 | -31.0685
20.8827 | -35.2324 | -31.0748
20.8888 | -35.2327 | -31.0811
20.8949 | -35.2331 | -31.0874
20.901 | -35.2336 | -31.0933
20.9071 | -35.234 | -31.0997
20.9132 | -35.2344 | -31.1061
20.9193 | -35.2349 | -31.1126
20.9254 | -35.2353 | -31.1189
20.9315 | -35.2359 | -31.1254
20.9376 | -35.2363 | -31.1321
20.9437 | -35.2367 | -31.1384
20.9498 | -35.2371 | -31.1447
20.9559 | -35.2376 | -31.1506
20.962 | -35.238 | -31.1575
20.9681 | -35.2385 | -31.1645
20.9742 | -35.239 | -31.1705
20.9803 | -35.2395 | -31.1763
20.9864 | -35.2399 | -31.1821
20.9926 | -35.2403 | -31.1889
20.9987 | -35.2408 | -31.1949
21.0048 | -35.2412 | -31.2011
21.0109 | -35.2418 | -31.2076
21.017 | -35.2423 | -31.2143
21.0231 | -35.2428 | -31.2206
21.0292 | -35.2432 | -31.2266
21.0353 | -35.2437 | -31.2328
21.0414 | -35.2441 | -31.239
21.0475 | -35.2446 | -31.2458
21.0536 | -35.2451 | -31.252
21.0597 | -35.2455 | -31.2581
21.0658 | -35.246 | -31.2649
21.0719 | -35.2465 | -31.2703
21.078 | -35.2469 | -31.2761
21.0841 | -35.2474 | -31.2822
21.0902 | -35.2479 | -31.2883
21.0963 | -35.2483 | -31.295
21.1024 | -35.2488 | -31.3016
21.1085 | -35.2493 | -31.3085
21.1146 | -35.2498 | -31.3148
21.1207 | -35.2503 | -31.3214
21.1268 | -35.2507 | -31.3276
21.133 | -35.2511 | -31.3339
21.1391 | -35.2516 | -31.3395
21.1452 | -35.2522 | -31.3453
21.1513 | -35.2527 | -31.3516
21.1574 | -35.2531 | -31.3578
21.1635 | -35.2537 | -31.3641
21.1696 | -35.254 | -31.37
21.1757 | -35.2545 | -31.3764
21.1818 | -35.255 | -31.3832
21.1879 | -35.2555 | -31.3894
21.194 | -35.256 | -31.3951
21.2001 | -35.2564 | -31.4017
21.2062 | -35.2568 | -31.4077
21.2123 | -35.2573 | -31.4143
21.2184 | -35.2578 | -31.4202
21.2245 | -35.2582 | -31.4259
21.2306 | -35.2587 | -31.432
21.2367 | -35.2591 | -31.4387
21.2428 | -35.2595 | -31.4451
21.2489 | -35.26 | -31.4517
21.255 | -35.2605 | -31.4579
21.2611 | -35.2609 | -31.4645
21.2672 | -35.2614 | -31.4709
21.2733 | -35.2619 | -31.4774
21.2795 | -35.2624 | -31.4838
21.2856 | -35.2629 | -31.4904
21.2917 | -35.2633 | -31.4965
21.2978 | -35.2638 | -31.5027
21.3039 | -35.2643 | -31.5094
21.31 | -35.2648 | -31.5157
21.3161 | -35.2652 | -31.5218
21.3222 | -35.2657 | -31.5282
21.3283 | -35.2661 | -31.534
21.3344 | -35.2666 | -31.5404
21.3405 | -35.267 | -31.5467
21.3466 | -35.2674 | -31.5529
21.3527 | -35.2681 | -31.5586
21.3588 | -35.2686 | -31.5647
21.3649 | -35.2691 | -31.5711
21.371 | -35.2694 | -31.5778
21.3771 | -35.2698 | -31.584
21.3832 | -35.2703 | -31.5904
21.3893 | -35.2707 | -31.5967
21.3954 | -35.2712 | -31.6029
21.4015 | -35.2717 | -31.6093
21.4076 | -35.2721 | -31.6155
21.4137 | -35.2725 | -31.622
21.4199 | -35.273 | -31.6284
21.426 | -35.2736 | -31.6328
21.4321 | -35.2741 | -31.639
21.4382 | -35.2746 | -31.6455
21.4443 | -35.275 | -31.6514
21.4504 | -35.2753 | -31.6574
21.4565 | -35.2758 | -31.6629
21.4626 | -35.2763 | -31.6689
21.4687 | -35.2768 | -31.6749
21.4748 | -35.2771 | -31.6807
21.4809 | -35.2777 | -31.6872
21.487 | -35.2781 | -31.6937
21.4931 | -35.2787 | -31.6994
21.4992 | -35.2791 | -31.7053
21.5053 | -35.2797 | -31.7119
21.5114 | -35.2802 | -31.719
21.5175 | -35.2807 | -31.725
21.5236 | -35.2812 | -31.7316
21.5297 | -35.2817 | -31.7365
21.5358 | -35.2823 | -31.7422
21.5419 | -35.2827 | -31.7487
21.548 | -35.2831 | -31.7547
21.5541 | -35.2835 | -31.7609
21.5602 | -35.284 | -31.7671
21.5664 | -35.2844 | -31.7721
21.5725 | -35.285 | -31.7779
21.5786 | -35.2854 | -31.7843
21.5847 | -35.2858 | -31.7905
21.5908 | -35.2864 | -31.797
21.5969 | -35.2869 | -31.8028
21.603 | -35.2875 | -31.8087
21.6091 | -35.2879 | -31.8155
21.6152 | -35.2884 | -31.8222
21.6213 | -35.2889 | -31.8287
21.6274 | -35.2894 | -31.8347
21.6335 | -35.2898 | -31.8408
21.6396 | -35.2901 | -31.8466
21.6457 | -35.2906 | -31.853
21.6518 | -35.2911 | -31.8585
21.6579 | -35.2916 | -31.8641
21.664 | -35.292 | -31.8704
21.6701 | -35.2923 | -31.8766
21.6762 | -35.2928 | -31.8822
21.6823 | -35.2933 | -31.8878
21.6884 | -35.2937 | -31.894
21.6945 | -35.2941 | -31.8999
21.7006 | -35.2945 | -31.906
21.7068 | -35.2952 | -31.9117
21.7129 | -35.2957 | -31.9185
21.719 | -35.2961 | -31.9249
21.7251 | -35.2966 | -31.9311
21.7312 | -35.297 | -31.9376
21.7373 | -35.2974 | -31.9428
21.7434 | -35.2978 | -31.9492
21.7495 | -35.2982 | -31.9553
21.7556 | -35.2986 | -31.961
21.7617 | -35.299 | -31.9668
21.7678 | -35.2994 | -31.9733
21.7739 | -35.2998 | -31.9784
21.78 | -35.3003 | -31.9846
21.7861 | -35.3007 | -31.9906
21.7922 | -35.3011 | -31.9966
21.7983 | -35.3015 | -32.0029
21.8044 | -35.302 | -32.0086
21.8105 | -35.3025 | -32.0148
21.8166 | -35.3031 | -32.0208
21.8227 | -35.3035 | -32.0275
21.8288 | -35.304 | -32.0339
21.8349 | -35.3045 | -32.0411
21.841 | -35.3049 | -32.0473
21.8471 | -35.3054 | -32.0531
21.8533 | -35.3059 | -32.0593
21.8594 | -35.3063 | -32.0651
21.8655 | -35.3068 | -32.0714
21.8716 | -35.3072 | -32.0772
21.8777 | -35.3076 | -32.0825
21.8838 | -35.3081 | -32.0879
21.8899 | -35.3085 | -32.0944
21.896 | -35.309 | -32.1008
21.9021 | -35.3095 | -32.1066
21.9082 | -35.3099 | -32.1121
21.9143 | -35.3104 | -32.1183
21.9204 | -35.3109 | -32.1248
21.9265 | -35.3114 | -32.131
21.9326 | -35.3118 | -32.1369
21.9387 | -35.3123 | -32.1427
21.9448 | -35.3128 | -32.1484
21.9509 | -35.3132 | -32.1548
21.957 | -35.3137 | -32.1607
21.9631 | -35.3141 | -32.1664
21.9692 | -35.3144 | -32.1724
21.9753 | -35.3149 | -32.1785
21.9814 | -35.3155 | -32.1844
21.9875 | -35.3158 | -32.1896
21.9937 | -35.3163 | -32.1957
21.9998 | -35.3168 | -32.2013
22.0059 | -35.3174 | -32.2077
22.012 | -35.3179 | -32.2134
22.0181 | -35.3182 | -32.2198
22.0242 | -35.3187 | -32.226
22.0303 | -35.3191 | -32.2326
22.0364 | -35.3195 | -32.2386
22.0425 | -35.32 | -32.2444
22.0486 | -35.3204 | -32.2514
22.0547 | -35.321 | -32.2567
22.0608 | -35.3214 | -32.262
22.0669 | -35.3219 | -32.268
22.073 | -35.3224 | -32.2735
22.0791 | -35.3229 | -32.2801
22.0852 | -35.3233 | -32.2863
22.0913 | -35.3238 | -32.2922
22.0974 | -35.3243 | -32.2976
22.1035 | -35.3247 | -32.3041
22.1096 | -35.3252 | -32.3108
22.1157 | -35.3258 | -32.3171
22.1218 | -35.3262 | -32.3229
22.1279 | -35.3268 | -32.328
22.134 | -35.3273 | -32.3337
22.1402 | -35.3278 | -32.34
22.1463 | -35.3284 | -32.3458
22.1524 | -35.3289 | -32.352
22.1585 | -35.3295 | -32.3578
22.1646 | -35.33 | -32.3629
22.1707 | -35.3305 | -32.3697
22.1768 | -35.331 | -32.3751
22.1829 | -35.3315 | -32.3804
22.189 | -35.3319 | -32.3863
22.1951 | -35.3323 | -32.3919
22.2012 | -35.3328 | -32.3979
22.2073 | -35.3333 | -32.4042
22.2134 | -35.3337 | -32.4098
22.2195 | -35.3341 | -32.4153
22.2256 | -35.3347 | -32.4217
22.2317 | -35.3351 | -32.427
22.2378 | -35.3356 | -32.4323
22.2439 | -35.3361 | -32.4383
22.25 | -35.3366 | -32.4445
22.2561 | -35.3369 | -32.4501
22.2622 | -35.3375 | -32.456
22.2683 | -35.3381 | -32.4611
22.2744 | -35.3385 | -32.4675
22.2806 | -35.3389 | -32.4736
22.2867 | -35.3394 | -32.479
22.2928 | -35.34 | -32.4849
22.2989 | -35.3405 | -32.4902
22.305 | -35.3409 | -32.4973
22.3111 | -35.3414 | -32.5039
22.3172 | -35.3419 | -32.5101
22.3233 | -35.3425 | -32.5161
22.3294 | -35.343 | -32.5223
22.3355 | -35.3434 | -32.5288
22.3416 | -35.3438 | -32.5356
22.3477 | -35.3444 | -32.5417
22.3538 | -35.3448 | -32.5474
22.3599 | -35.3453 | -32.5534
22.366 | -35.3457 | -32.5589
22.3721 | -35.3463 | -32.5648
22.3782 | -35.3467 | -32.5705
22.3843 | -35.3473 | -32.5767
22.3904 | -35.3477 | -32.5826
22.3965 | -35.3482 | -32.5883
22.4026 | -35.3486 | -32.5941
22.4087 | -35.3491 | -32.6001
22.4148 | -35.3496 | -32.6065
22.4209 | -35.35 | -32.6123
22.4271 | -35.3506 | -32.6184
22.4332 | -35.3511 | -32.6246
22.4393 | -35.3515 | -32.6307
22.4454 | -35.352 | -32.6366
22.4515 | -35.3525 | -32.6421
22.4576 | -35.3529 | -32.6487
22.4637 | -35.3534 | -32.655
22.4698 | -35.3539 | -32.6611
22.4759 | -35.3545 | -32.6678
22.482 | -35.355 | -32.6736
22.4881 | -35.3555 | -32.6795
22.4942 | -35.356 | -32.6853
22.5003 | -35.3566 | -32.6913
22.5064 | -35.3572 | -32.6975
22.5125 | -35.3577 | -32.7024
22.5186 | -35.3582 | -32.7089
22.5247 | -35.3585 | -32.7143
22.5308 | -35.3589 | -32.7203
22.5369 | -35.3593 | -32.7266
22.543 | -35.3598 | -32.7331
22.5491 | -35.3602 | -32.7389
22.5552 | -35.3607 | -32.7447
22.5613 | -35.3611 | -32.7507
22.5675 | -35.3615 | -32.7566
22.5736 | -35.3619 | -32.7628
22.5797 | -35.3625 | -32.7701
22.5858 | -35.3629 | -32.7756
22.5919 | -35.3635 | -32.781
22.598 | -35.364 | -32.7871
22.6041 | -35.3645 | -32.793
22.6102 | -35.3649 | -32.7992
22.6163 | -35.3655 | -32.8055
22.6224 | -35.3661 | -32.8119
22.6285 | -35.3666 | -32.8171
22.6346 | -35.367 | -32.8234
22.6407 | -35.3675 | -32.8299
22.6468 | -35.368 | -32.8353
22.6529 | -35.3685 | -32.8409
22.659 | -35.3689 | -32.8459
22.6651 | -35.3695 | -32.8516
22.6712 | -35.37 | -32.858
22.6773 | -35.3704 | -32.8641
22.6834 | -35.3708 | -32.8702
22.6895 | -35.3714 | -32.8758
22.6956 | -35.3719 | -32.8812
22.7017 | -35.3724 | -32.8872
22.7079 | -35.373 | -32.8929
22.714 | -35.3735 | -32.8986
22.7201 | -35.374 | -32.9043
22.7262 | -35.3747 | -32.9092
22.7323 | -35.3751 | -32.9155
22.7384 | -35.3756 | -32.9208
22.7445 | -35.3761 | -32.9269
22.7506 | -35.3765 | -32.9326
22.7567 | -35.3769 | -32.9384
22.7628 | -35.3774 | -32.9449
22.7689 | -35.3779 | -32.9504
22.775 | -35.3784 | -32.9562
22.7811 | -35.3789 | -32.9612
22.7872 | -35.3793 | -32.9672
22.7933 | -35.38 | -32.9732
22.7994 | -35.3805 | -32.9792
22.8055 | -35.3811 | -32.9851
22.8116 | -35.3816 | -32.9908
22.8177 | -35.3821 | -32.9969
22.8238 | -35.3826 | -33.0019
22.8299 | -35.3831 | -33.0072
22.836 | -35.3836 | -33.0138
22.8421 | -35.3841 | -33.0198
22.8482 | -35.3846 | -33.0258
22.8544 | -35.385 | -33.0312
22.8605 | -35.3854 | -33.0365
22.8666 | -35.386 | -33.0434
22.8727 | -35.3864 | -33.0493
22.8788 | -35.3868 | -33.0551
22.8849 | -35.3874 | -33.061
22.891 | -35.3879 | -33.0672
22.8971 | -35.3884 | -33.0729
22.9032 | -35.3889 | -33.0776
22.9093 | -35.3894 | -33.0832
22.9154 | -35.3899 | -33.0894
22.9215 | -35.3903 | -33.0952
22.9276 | -35.3907 | -33.101
22.9337 | -35.3912 | -33.1061
22.9398 | -35.3916 | -33.1116
22.9459 | -35.3921 | -33.1176
22.952 | -35.3925 | -33.1234
22.9581 | -35.3929 | -33.1299
22.9642 | -35.3933 | -33.1358
22.9703 | -35.3938 | -33.1423
22.9764 | -35.3944 | -33.1484
22.9825 | -35.3949 | -33.1542
22.9886 | -35.3954 | -33.1601
22.9948 | -35.396 | -33.1656
23.0009 | -35.3964 | -33.1715
23.007 | -35.3969 | -33.1775
23.0131 | -35.3974 | -33.1831
23.0192 | -35.3978 | -33.1886
23.0253 | -35.3983 | -33.195
23.0314 | -35.3987 | -33.2009
23.0375 | -35.3993 | -33.2072
23.0436 | -35.3998 | -33.2118
23.0497 | -35.4004 | -33.2173
23.0558 | -35.4009 | -33.2237
23.0619 | -35.4013 | -33.2295
23.068 | -35.4018 | -33.2348
23.0741 | -35.4024 | -33.2389
23.0802 | -35.4028 | -33.2449
23.0863 | -35.4033 | -33.2508
23.0924 | -35.4037 | -33.2569
23.0985 | -35.4041 | -33.2627
23.1046 | -35.4046 | -33.2684
23.1107 | -35.4051 | -33.2734
23.1168 | -35.4054 | -33.2795
23.1229 | -35.406 | -33.2857
23.129 | -35.4065 | -33.2921
23.1351 | -35.4071 | -33.2982
23.1413 | -35.4076 | -33.3034
23.1474 | -35.4081 | -33.3087
23.1535 | -35.4085 | -33.3146
23.1596 | -35.4091 | -33.3206
23.1657 | -35.4096 | -33.3267
23.1718 | -35.4102 | -33.333
23.1779 | -35.4108 | -33.3394
23.184 | -35.4114 | -33.3457
23.1901 | -35.4119 | -33.3515
23.1962 | -35.4123 | -33.3573
23.2023 | -35.4129 | -33.364
23.2084 | -35.4134 | -33.3695
23.2145 | -35.4139 | -33.3739
23.2206 | -35.4143 | -33.3793
23.2267 | -35.4148 | -33.3847
23.2328 | -35.4152 | -33.3914
23.2389 | -35.4157 | -33.397
23.245 | -35.4161 | -33.403
23.2511 | -35.4166 | -33.4078
23.2572 | -35.417 | -33.4135
23.2633 | -35.4175 | -33.4191
23.2694 | -35.418 | -33.4248
23.2755 | -35.4184 | -33.4307
23.2817 | -35.419 | -33.4368
23.2878 | -35.4193 | -33.4424
23.2939 | -35.4197 | -33.4481
23.3 | -35.4202 | -33.4543
23.3061 | -35.4206 | -33.4605
23.3122 | -35.4213 | -33.4663
23.3183 | -35.4218 | -33.4721
23.3244 | -35.4222 | -33.477
23.3305 | -35.4226 | -33.4831
23.3366 | -35.4231 | -33.4886
23.3427 | -35.4236 | -33.494
23.3488 | -35.4241 | -33.4999
23.3549 | -35.4246 | -33.5058
23.361 | -35.4252 | -33.5122
23.3671 | -35.4257 | -33.5183
23.3732 | -35.426 | -33.5245
23.3793 | -35.4265 | -33.5296
23.3854 | -35.4271 | -33.5359
23.3915 | -35.4276 | -33.5412
23.3976 | -35.428 | -33.5477
23.4037 | -35.4285 | -33.5525
23.4098 | -35.4289 | -33.5588
23.4159 | -35.4295 | -33.5647
23.422 | -35.43 | -33.5705
23.4282 | -35.4306 | -33.5766
23.4343 | -35.431 | -33.5827
23.4404 | -35.4316 | -33.5886
23.4465 | -35.4321 | -33.5941
23.4526 | -35.4326 | -33.6001
23.4587 | -35.4331 | -33.6051
23.4648 | -35.4337 | -33.6108
23.4709 | -35.434 | -33.6167
23.477 | -35.4345 | -33.6223
23.4831 | -35.4349 | -33.6283
23.4892 | -35.4353 | -33.634
23.4953 | -35.4358 | -33.6405
23.5014 | -35.4362 | -33.646
23.5075 | -35.4367 | -33.6513
23.5136 | -35.4371 | -33.6577
23.5197 | -35.4376 | -33.6634
23.5258 | -35.438 | -33.6698
23.5319 | -35.4385 | -33.6754
23.538 | -35.439 | -33.6815
23.5441 | -35.4396 | -33.6868
23.5502 | -35.4401 | -33.6925
23.5563 | -35.4405 | -33.6993
23.5624 | -35.441 | -33.7053
23.5686 | -35.4414 | -33.7116
23.5747 | -35.4419 | -33.7171
23.5808 | -35.4423 | -33.7221
23.5869 | -35.4427 | -33.7281
23.593 | -35.4432 | -33.7346
23.5991 | -35.4436 | -33.7402
23.6052 | -35.4441 | -33.7455
23.6113 | -35.4447 | -33.752
23.6174 | -35.4453 | -33.7582
23.6235 | -35.4457 | -33.7646
23.6296 | -35.4461 | -33.7709
23.6357 | -35.4465 | -33.7761
23.6418 | -35.4471 | -33.7826
23.6479 | -35.4476 | -33.7877
23.654 | -35.4482 | -33.7933
23.6601 | -35.4487 | -33.799
23.6662 | -35.4491 | -33.805
23.6723 | -35.4495 | -33.8114
23.6784 | -35.4498 | -33.817
23.6845 | -35.4502 | -33.8226
23.6906 | -35.4507 | -33.8278
23.6967 | -35.4511 | -33.8339
23.7028 | -35.4516 | -33.8404
23.7089 | -35.4519 | -33.8462
23.7151 | -35.4524 | -33.852
23.7212 | -35.4527 | -33.8577
23.7273 | -35.4532 | -33.8641
23.7334 | -35.4538 | -33.8699
23.7395 | -35.4542 | -33.8762
23.7456 | -35.4548 | -33.8816
23.7517 | -35.4552 | -33.8868
23.7578 | -35.4556 | -33.892
23.7639 | -35.4561 | -33.898
23.77 | -35.4565 | -33.9037
23.7761 | -35.457 | -33.9096
23.7822 | -35.4576 | -33.9154
23.7883 | -35.458 | -33.9215
23.7944 | -35.4585 | -33.9278
23.8005 | -35.4589 | -33.9341
23.8066 | -35.4593 | -33.9405
23.8127 | -35.4597 | -33.9461
23.8188 | -35.4602 | -33.9524
23.8249 | -35.4607 | -33.9586
23.831 | -35.4612 | -33.9645
23.8371 | -35.4617 | -33.9698
23.8432 | -35.4621 | -33.975
23.8493 | -35.4626 | -33.9806
23.8555 | -35.4631 | -33.9866
23.8616 | -35.4635 | -33.9928
23.8677 | -35.464 | -33.9985
23.8738 | -35.4644 | -34.004
23.8799 | -35.4648 | -34.0096
23.886 | -35.4656 | -34.0157
23.8921 | -35.466 | -34.0214
23.8982 | -35.4665 | -34.027
23.9043 | -35.4669 | -34.0327
23.9104 | -35.4674 | -34.0388
23.9165 | -35.4677 | -34.0443
23.9226 | -35.468 | -34.0507
23.9287 | -35.4685 | -34.0565
23.9348 | -35.4689 | -34.0623
23.9409 | -35.4693 | -34.0679
23.947 | -35.4698 | -34.0737
23.9531 | -35.4703 | -34.079
23.9592 | -35.4706 | -34.0843
23.9653 | -35.4712 | -34.0911
23.9714 | -35.4717 | -34.0962
23.9775 | -35.4722 | -34.1016
23.9836 | -35.4726 | -34.1067
23.9897 | -35.4729 | -34.1126
23.9958 | -35.4734 | -34.1182
24.002 | -35.4739 | -34.1238
24.0081 | -35.4744 | -34.1301
24.0142 | -35.4749 | -34.136
24.0203 | -35.4755 | -34.1421
24.0264 | -35.4759 | -34.1484
24.0325 | -35.4763 | -34.1541
24.0386 | -35.4767 | -34.1602
24.0447 | -35.4772 | -34.1663
24.0508 | -35.4776 | -34.1725
24.0569 | -35.478 | -34.1786
24.063 | -35.4784 | -34.1839
24.0691 | -35.4788 | -34.1893
24.0752 | -35.4793 | -34.1954
24.0813 | -35.4798 | -34.2007
24.0874 | -35.4803 | -34.2064
24.0935 | -35.4808 | -34.2124
24.0996 | -35.4813 | -34.2188
24.1057 | -35.4818 | -34.2245
24.1118 | -35.4823 | -34.2306
24.1179 | -35.4827 | -34.2369
24.124 | -35.483 | -34.2433
24.1301 | -35.4836 | -34.2489
24.1362 | -35.4841 | -34.2551
24.1424 | -35.4845 | -34.2615
24.1485 | -35.4849 | -34.2671
24.1546 | -35.4853 | -34.273
24.1607 | -35.4857 | -34.2791
24.1668 | -35.4861 | -34.2859
24.1729 | -35.4866 | -34.2922
24.179 | -35.4872 | -34.2983
24.1851 | -35.4876 | -34.3035
24.1912 | -35.4882 | -34.3097
24.1973 | -35.4886 | -34.3154
24.2034 | -35.4891 | -34.3204
24.2095 | -35.4896 | -34.3265
24.2156 | -35.4901 | -34.3328
24.2217 | -35.4905 | -34.3384
24.2278 | -35.491 | -34.3441
24.2339 | -35.4914 | -34.3491
24.24 | -35.4918 | -34.3565
24.2461 | -35.4922 | -34.3615
24.2522 | -35.4926 | -34.3669
24.2583 | -35.493 | -34.3726
24.2644 | -35.4934 | -34.3789
24.2705 | -35.4938 | -34.385
24.2766 | -35.4943 | -34.3916
24.2827 | -35.4949 | -34.3975
24.2889 | -35.4954 | -34.4032
24.295 | -35.4959 | -34.4085
24.3011 | -35.4963 | -34.4138
24.3072 | -35.4968 | -34.4197
24.3133 | -35.4972 | -34.4255
24.3194 | -35.4975 | -34.4315
24.3255 | -35.4978 | -34.4374
24.3316 | -35.4983 | -34.4435
24.3377 | -35.4987 | -34.4498
24.3438 | -35.499 | -34.4554
24.3499 | -35.4995 | -34.4611
24.356 | -35.5 | -34.4667
24.3621 | -35.5004 | -34.4727
24.3682 | -35.5008 | -34.4785
24.3743 | -35.5012 | -34.484
24.3804 | -35.5016 | -34.4894
24.3865 | -35.502 | -34.4954
24.3926 | -35.5024 | -34.5006
24.3987 | -35.5028 | -34.5065
24.4048 | -35.5032 | -34.5123
24.4109 | -35.5036 | -34.5176
24.417 | -35.504 | -34.5233
24.4231 | -35.5045 | -34.5293
24.4293 | -35.5049 | -34.5347
24.4354 | -35.5054 | -34.5408
24.4415 | -35.5059 | -34.5465
24.4476 | -35.5063 | -34.5527
24.4537 | -35.5066 | -34.5583
24.4598 | -35.5071 | -34.5636
24.4659 | -35.5075 | -34.5696
24.472 | -35.508 | -34.5753
24.4781 | -35.5084 | -34.5806
24.4842 | -35.5089 | -34.5875
24.4903 | -35.5093 | -34.5932
24.4964 | -35.5096 | -34.5991
24.5025 | -35.51 | -34.6054
24.5086 | -35.5105 | -34.6111
24.5147 | -35.5109 | -34.6173
24.5208 | -35.5113 | -34.623
24.5269 | -35.5118 | -34.6285
24.533 | -35.5121 | -34.6346
24.5391 | -35.5126 | -34.6395
24.5452 | -35.513 | -34.6455
24.5513 | -35.5133 | -34.6509
24.5574 | -35.5137 | -34.6567
24.5635 | -35.5142 | -34.6626
24.5696 | -35.5146 | -34.6688
24.5758 | -35.515 | -34.6741
24.5819 | -35.5155 | -34.6793
24.588 | -35.5157 | -34.6852
24.5941 | -35.5161 | -34.6911
24.6002 | -35.5166 | -34.6974
24.6063 | -35.5171 | -34.7028
24.6124 | -35.5176 | -34.7089
24.6185 | -35.5181 | -34.7149
24.6246 | -35.5185 | -34.7208
24.6307 | -35.5189 | -34.7263
24.6368 | -35.5194 | -34.7321
24.6429 | -35.5197 | -34.739
24.649 | -35.5201 | -34.7448
24.6551 | -35.5206 | -34.7504
24.6612 | -35.521 | -34.7561
24.6673 | -35.5213 | -34.7619
24.6734 | -35.5217 | -34.7681
24.6795 | -35.5221 | -34.7737
24.6856 | -35.5226 | -34.7787
24.6917 | -35.523 | -34.784
24.6978 | -35.5233 | -34.7901
24.7039 | -35.5238 | -34.7955
24.71 | -35.5241 | -34.8013
24.7162 | -35.5244 | -34.8073
24.7223 | -35.5249 | -34.8131
24.7284 | -35.5253 | -34.8186
24.7345 | -35.5257 | -34.8238
24.7406 | -35.5261 | -34.8297
24.7467 | -35.5265 | -34.8351
24.7528 | -35.5269 | -34.8413
24.7589 | -35.5273 | -34.8469
24.765 | -35.5277 | -34.852
24.7711 | -35.5281 | -34.8576
24.7772 | -35.5285 | -34.8634
24.7833 | -35.5289 | -34.8689
24.7894 | -35.5293 | -34.8742
24.7955 | -35.5297 | -34.8802
24.8016 | -35.53 | -34.8862
24.8077 | -35.5304 | -34.8918
24.8138 | -35.5309 | -34.8977
24.8199 | -35.5313 | -34.9038
24.826 | -35.5317 | -34.9098
24.8321 | -35.5322 | -34.9156
24.8382 | -35.5326 | -34.9218
24.8443 | -35.5331 | -34.928
24.8504 | -35.5334 | -34.9336
24.8565 | -35.5339 | -34.9398
24.8627 | -35.5343 | -34.9456
24.8688 | -35.5346 | -34.9508
24.8749 | -35.535 | -34.9563
24.881 | -35.5354 | -34.9619
24.8871 | -35.5358 | -34.9681
24.8932 | -35.5363 | -34.974
24.8993 | -35.5368 | -34.9795
24.9054 | -35.5372 | -34.9855
24.9115 | -35.5377 | -34.9913
24.9176 | -35.538 | -34.9969
24.9237 | -35.5385 | -35.0024
24.9298 | -35.5388 | -35.0081
24.9359 | -35.5393 | -35.0145
24.942 | -35.5395 | -35.0204
24.9481 | -35.5399 | -35.0263
24.9542 | -35.5403 | -35.0324
24.9603 | -35.5407 | -35.038
24.9664 | -35.5411 | -35.0441
24.9725 | -35.5415 | -35.0502
24.9786 | -35.542 | -35.056
24.9847 | -35.5424 | -35.0613
24.9908 | -35.5428 | -35.0667
24.9969 | -35.5432 | -35.0723
25.0031 | -35.5435 | -35.0785
25.0092 | -35.5439 | -35.0841
25.0153 | -35.5443 | -35.0901
25.0214 | -35.5447 | -35.0955
25.0275 | -35.5451 | -35.101
25.0336 | -35.5456 | -35.1066
25.0397 | -35.5459 | -35.1126
25.0458 | -35.5462 | -35.1188
25.0519 | -35.5467 | -35.1246
25.058 | -35.5471 | -35.1305
25.0641 | -35.5475 | -35.136
25.0702 | -35.5478 | -35.1416
25.0763 | -35.5482 | -35.1477
25.0824 | -35.5486 | -35.1536
25.0885 | -35.549 | -35.159
25.0946 | -35.5494 | -35.1648
25.1007 | -35.5498 | -35.1703
25.1068 | -35.5501 | -35.1757
25.1129 | -35.5506 | -35.1817
25.119 | -35.5509 | -35.1873
25.1251 | -35.5513 | -35.1931
25.1312 | -35.5516 | -35.1989
25.1373 | -35.552 | -35.2048
25.1435 | -35.5525 | -35.2106
25.1496 | -35.5528 | -35.2159
25.1557 | -35.5532 | -35.2219
25.1618 | -35.5535 | -35.2278
25.1679 | -35.5539 | -35.2336
25.174 | -35.5543 | -35.2395
25.1801 | -35.5547 | -35.2455
25.1862 | -35.5549 | -35.2513
25.1923 | -35.5552 | -35.257
25.1984 | -35.5556 | -35.2629
25.2045 | -35.556 | -35.2683
25.2106 | -35.5564 | -35.2741
25.2167 | -35.5567 | -35.2801
25.2228 | -35.557 | -35.2858
25.2289 | -35.5574 | -35.2914
25.235 | -35.5578 | -35.2977
25.2411 | -35.5582 | -35.3034
25.2472 | -35.5585 | -35.3091
25.2533 | -35.5589 | -35.3149
25.2594 | -35.5593 | -35.3199
25.2655 | -35.5597 | -35.3255
25.2716 | -35.5601 | -35.3313
25.2777 | -35.5605 | -35.3367
25.2838 | -35.5609 | -35.3423
25.29 | -35.5613 | -35.3479
25.2961 | -35.5616 | -35.3535
25.3022 | -35.562 | -35.3591
25.3083 | -35.5624 | -35.3649
25.3144 | -35.5628 | -35.3706
25.3205 | -35.5632 | -35.3763
25.3266 | -35.5636 | -35.382
25.3327 | -35.564 | -35.3872
25.3388 | -35.5644 | -35.3929
25.3449 | -35.5648 | -35.3983
25.351 | -35.5651 | -35.4038
25.3571 | -35.5655 | -35.4098
25.3632 | -35.5658 | -35.4154
25.3693 | -35.5663 | -35.4212
25.3754 | -35.5667 | -35.4268
25.3815 | -35.5671 | -35.4328
25.3876 | -35.5674 | -35.4384
25.3937 | -35.5678 | -35.4442
25.3998 | -35.5681 | -35.4494
25.4059 | -35.5685 | -35.4551
25.412 | -35.5689 | -35.4609
25.4181 | -35.5692 | -35.4667
25.4242 | -35.5697 | -35.472
25.4304 | -35.57 | -35.4779
25.4365 | -35.5703 | -35.4835
25.4426 | -35.5707 | -35.489
25.4487 | -35.5711 | -35.4947
25.4548 | -35.5715 | -35.5005
25.4609 | -35.5718 | -35.5065
25.467 | -35.5722 | -35.5123
25.4731 | -35.5725 | -35.5182
25.4792 | -35.5729 | -35.5235
25.4853 | -35.5732 | -35.5294
25.4914 | -35.5737 | -35.5351
25.4975 | -35.574 | -35.5407
25.5036 | -35.5743 | -35.5465
25.5097 | -35.5747 | -35.5523
25.5158 | -35.575 | -35.5584
25.5219 | -35.5754 | -35.5644
25.528 | -35.5758 | -35.5699
25.5341 | -35.5762 | -35.5757
25.5402 | -35.5765 | -35.5816
25.5463 | -35.5768 | -35.5874
25.5524 | -35.5772 | -35.5931
25.5585 | -35.5776 | -35.599
25.5646 | -35.5779 | -35.6044
25.5707 | -35.5783 | -35.6104
25.5769 | -35.5785 | -35.6156
25.583 | -35.5788 | -35.6216
25.5891 | -35.5793 | -35.6275
25.5952 | -35.5796 | -35.6333
25.6013 | -35.5801 | -35.639
25.6074 | -35.5804 | -35.6449
25.6135 | -35.5808 | -35.6506
25.6196 | -35.5811 | -35.6561
25.6257 | -35.5815 | -35.662
25.6318 | -35.5819 | -35.6678
25.6379 | -35.5822 | -35.6731
25.644 | -35.5826 | -35.6782
25.6501 | -35.5829 | -35.6842
25.6562 | -35.5833 | -35.6901
25.6623 | -35.5837 | -35.6956
25.6684 | -35.584 | -35.7013
25.6745 | -35.5844 | -35.7074
25.6806 | -35.5847 | -35.7127
25.6867 | -35.585 | -35.7183
25.6928 | -35.5854 | -35.7239
25.6989 | -35.5857 | -35.73
25.705 | -35.5861 | -35.7354
25.7111 | -35.5864 | -35.7417
25.7173 | -35.5867 | -35.7473
25.7234 | -35.587 | -35.753
25.7295 | -35.5874 | -35.7588
25.7356 | -35.5878 | -35.7645
25.7417 | -35.5881 | -35.7703
25.7478 | -35.5883 | -35.7757
25.7539 | -35.5886 | -35.7811
25.76 | -35.589 | -35.7866
25.7661 | -35.5894 | -35.7923
25.7722 | -35.5897 | -35.7979
25.7783 | -35.5901 | -35.8033
25.7844 | -35.5905 | -35.8088
25.7905 | -35.5909 | -35.8143
25.7966 | -35.5912 | -35.82
25.8027 | -35.5916 | -35.8257
25.8088 | -35.592 | -35.8315
25.8149 | -35.5923 | -35.8372
25.821 | -35.5926 | -35.8425
25.8271 | -35.5929 | -35.8478
25.8332 | -35.5933 | -35.8532
25.8393 | -35.5936 | -35.8589
25.8454 | -35.5939 | -35.8647
25.8515 | -35.5942 | -35.8702
25.8576 | -35.5946 | -35.8755
25.8638 | -35.5949 | -35.881
25.8699 | -35.5952 | -35.8867
25.876 | -35.5956 | -35.8923
25.8821 | -35.5959 | -35.8983
25.8882 | -35.5962 | -35.9037
25.8943 | -35.5965 | -35.9094
25.9004 | -35.5968 | -35.9152
25.9065 | -35.5972 | -35.9212
25.9126 | -35.5975 | -35.9266
25.9187 | -35.5979 | -35.9321
25.9248 | -35.5981 | -35.9375
25.9309 | -35.5984 | -35.9434
25.937 | -35.5988 | -35.9491
25.9431 | -35.5991 | -35.9547
25.9492 | -35.5994 | -35.9602
25.9553 | -35.5998 | -35.966
25.9614 | -35.6 | -35.9713
25.9675 | -35.6004 | -35.9769
25.9736 | -35.6008 | -35.9827
25.9797 | -35.6011 | -35.9883
25.9858 | -35.6015 | -35.9938
25.9919 | -35.6018 | -35.9993
25.998 | -35.6022 | -36.0046
26.0042 | -35.6025 | -36.0101
26.0103 | -35.6028 | -36.0157
26.0164 | -35.6032 | -36.0211
26.0225 | -35.6035 | -36.0267
26.0286 | -35.6039 | -36.0322
26.0347 | -35.6041 | -36.0373
26.0408 | -35.6044 | -36.0429
26.0469 | -35.6047 | -36.0484
26.053 | -35.605 | -36.0543
26.0591 | -35.6053 | -36.0597
26.0652 | -35.6056 | -36.0651
26.0713 | -35.6059 | -36.0708
26.0774 | -35.6062 | -36.0764
26.0835 | -35.6065 | -36.0818
26.0896 | -35.6068 | -36.0873
26.0957 | -35.6071 | -36.0931
26.1018 | -35.6074 | -36.0983
26.1079 | -35.6078 | -36.1037
26.114 | -35.6081 | -36.1095
26.1201 | -35.6084 | -36.1152
26.1262 | -35.6088 | -36.1209
26.1323 | -35.609 | -36.126
26.1384 | -35.6094 | -36.1314
26.1445 | -35.6096 | -36.1372
26.1507 | -35.61 | -36.1425
26.1568 | -35.6103 | -36.1482
26.1629 | -35.6106 | -36.1533
26.169 | -35.6109 | -36.1589
26.1751 | -35.6111 | -36.1644
26.1812 | -35.6115 | -36.1701
26.1873 | -35.6118 | -36.1755
26.1934 | -35.6121 | -36.181
26.1995 | -35.6125 | -36.1867
26.2056 | -35.6128 | -36.1921
26.2117 | -35.6131 | -36.198
26.2178 | -35.6134 | -36.2033
26.2239 | -35.6137 | -36.2086
26.23 | -35.614 | -36.2142
26.2361 | -35.6144 | -36.2195
26.2422 | -35.6147 | -36.2247
26.2483 | -35.6149 | -36.2304
26.2544 | -35.6152 | -36.236
26.2605 | -35.6156 | -36.2419
26.2666 | -35.6159 | -36.2474
26.2727 | -35.6162 | -36.2527
26.2788 | -35.6165 | -36.258
26.2849 | -35.6169 | -36.2633
26.2911 | -35.6172 | -36.2683
26.2972 | -35.6175 | -36.274
26.3033 | -35.6178 | -36.2797
26.3094 | -35.6182 | -36.2853
26.3155 | -35.6184 | -36.291
26.3216 | -35.6187 | -36.2966
26.3277 | -35.619 | -36.302
26.3338 | -35.6192 | -36.3074
26.3399 | -35.6195 | -36.3128
26.346 | -35.6199 | -36.3181
26.3521 | -35.6202 | -36.3237
26.3582 | -35.6206 | -36.3294
26.3643 | -35.621 | -36.3349
26.3704 | -35.6213 | -36.3406
26.3765 | -35.6216 | -36.3464
26.3826 | -35.6218 | -36.3522
26.3887 | -35.622 | -36.3575
26.3948 | -35.6224 | -36.3627
26.4009 | -35.6227 | -36.3681
26.407 | -35.623 | -36.3733
26.4131 | -35.6233 | -36.3788
26.4192 | -35.6236 | -36.3837
26.4253 | -35.6239 | -36.3894
26.4314 | -35.6243 | -36.3947
26.4376 | -35.6246 | -36.4003
26.4437 | -35.6249 | -36.4057
26.4498 | -35.6252 | -36.4114
26.4559 | -35.6255 | -36.4164
26.462 | -35.6258 | -36.4217
26.4681 | -35.6261 | -36.427
26.4742 | -35.6264 | -36.4326
26.4803 | -35.6267 | -36.4383
26.4864 | -35.627 | -36.4435
26.4925 | -35.6274 | -36.4488
26.4986 | -35.6276 | -36.4541
26.5047 | -35.628 | -36.46
26.5108 | -35.6283 | -36.4652
26.5169 | -35.6286 | -36.4705
26.523 | -35.6288 | -36.4757
26.5291 | -35.6291 | -36.4811
26.5352 | -35.6294 | -36.4865
26.5413 | -35.6297 | -36.4919
26.5474 | -35.63 | -36.4976
26.5535 | -35.6303 | -36.503
26.5596 | -35.6306 | -36.5083
26.5657 | -35.6309 | -36.5139
26.5718 | -35.6312 | -36.5194
26.578 | -35.6315 | -36.525
26.5841 | -35.6318 | -36.5303
26.5902 | -35.6321 | -36.5358
26.5963 | -35.6323 | -36.5411
26.6024 | -35.6326 | -36.5467
26.6085 | -35.6329 | -36.5519
26.6146 | -35.6332 | -36.5572
26.6207 | -35.6335 | -36.5624
26.6268 | -35.6338 | -36.568
26.6329 | -35.6341 | -36.5731
26.639 | -35.6344 | -36.578
26.6451 | -35.6347 | -36.5833
26.6512 | -35.635 | -36.5883
26.6573 | -35.6352 | -36.5941
26.6634 | -35.6354 | -36.5991
26.6695 | -35.6358 | -36.6045
26.6756 | -35.636 | -36.6101
26.6817 | -35.6363 | -36.6151
26.6878 | -35.6366 | -36.6209
26.6939 | -35.6369 | -36.6263
26.7 | -35.6372 | -36.6314
26.7061 | -35.6375 | -36.6366
26.7122 | -35.6378 | -36.642
26.7183 | -35.638 | -36.6474
26.7245 | -35.6383 | -36.6526
26.7306 | -35.6387 | -36.6577
26.7367 | -35.6389 | -36.6629
26.7428 | -35.6392 | -36.668
26.7489 | -35.6395 | -36.6734
26.755 | -35.6398 | -36.6787
26.7611 | -35.6402 | -36.684
26.7672 | -35.6405 | -36.6898
26.7733 | -35.6407 | -36.6949
26.7794 | -35.641 | -36.7001
26.7855 | -35.6413 | -36.7058
26.7916 | -35.6416 | -36.7106
26.7977 | -35.6418 | -36.7159
26.8038 | -35.642 | -36.7212
26.8099 | -35.6424 | -36.7264
26.816 | -35.6426 | -36.7321
26.8221 | -35.6428 | -36.7373
26.8282 | -35.6431 | -36.7423
26.8343 | -35.6434 | -36.7478
26.8404 | -35.6437 | -36.7529
26.8465 | -35.644 | -36.7583
26.8526 | -35.6442 | -36.7637
26.8587 | -35.6445 | -36.7692
26.8649 | -35.6448 | -36.7745
26.871 | -35.6451 | -36.7797
26.8771 | -35.6454 | -36.7846
26.8832 | -35.6457 | -36.7902
26.8893 | -35.646 | -36.7955
26.8954 | -35.6462 | -36.801
26.9015 | -35.6465 | -36.8061
26.9076 | -35.6467 | -36.8117
26.9137 | -35.647 | -36.8166
26.9198 | -35.6472 | -36.8223
26.9259 | -35.6475 | -36.8274
26.932 | -35.6478 | -36.833
26.9381 | -35.6481 | -36.8385
26.9442 | -35.6484 | -36.8436
26.9503 | -35.6487 | -36.8491
26.9564 | -35.6489 | -36.8545
26.9625 | -35.6492 | -36.8598
26.9686 | -35.6495 | -36.865
26.9747 | -35.6498 | -36.8703
26.9808 | -35.6501 | -36.8754
26.9869 | -35.6503 | -36.8805
26.993 | -35.6506 | -36.8859
26.9991 | -35.6508 | -36.8912
27.0052 | -35.6511 | -36.8964
27.0114 | -35.6514 | -36.9019
27.0175 | -35.6516 | -36.907
27.0236 | -35.652 | -36.9127
27.0297 | -35.6522 | -36.918
27.0358 | -35.6525 | -36.9235
27.0419 | -35.6528 | -36.9291
27.048 | -35.6531 | -36.9342
27.0541 | -35.6533 | -36.9394
27.0602 | -35.6536 | -36.945
27.0663 | -35.6538 | -36.9505
27.0724 | -35.654 | -36.9561
27.0785 | -35.6544 | -36.9616
27.0846 | -35.6546 | -36.9672
27.0907 | -35.6549 | -36.9724
27.0968 | -35.6552 | -36.9778
27.1029 | -35.6554 | -36.9831
27.109 | -35.6557 | -36.9884
27.1151 | -35.656 | -36.9938
27.1212 | -35.6562 | -36.9989
27.1273 | -35.6565 | -37.0043
27.1334 | -35.6568 | -37.0097
27.1395 | -35.657 | -37.015
27.1456 | -35.6573 | -37.0204
27.1518 | -35.6576 | -37.0258
27.1579 | -35.6578 | -37.031
27.164 | -35.6581 | -37.0361
27.1701 | -35.6584 | -37.0413
27.1762 | -35.6587 | -37.0464
27.1823 | -35.6589 | -37.0519
27.1884 | -35.6592 | -37.0572
27.1945 | -35.6594 | -37.062
27.2006 | -35.6597 | -37.0669
27.2067 | -35.66 | -37.0723
27.2128 | -35.6603 | -37.0776
27.2189 | -35.6606 | -37.0828
27.225 | -35.6609 | -37.0879
27.2311 | -35.6612 | -37.0932
27.2372 | -35.6614 | -37.0984
27.2433 | -35.6617 | -37.1033
27.2494 | -35.6621 | -37.1086
27.2555 | -35.6623 | -37.114
27.2616 | -35.6625 | -37.119
27.2677 | -35.6627 | -37.1246
27.2738 | -35.663 | -37.1296
27.2799 | -35.6633 | -37.135
27.286 | -35.6635 | -37.14
27.2921 | -35.6638 | -37.1454
27.2983 | -35.6641 | -37.1501
27.3044 | -35.6644 | -37.1556
27.3105 | -35.6646 | -37.1609
27.3166 | -35.6649 | -37.1663
27.3227 | -35.6651 | -37.1713
27.3288 | -35.6654 | -37.1767
27.3349 | -35.6656 | -37.1818
27.341 | -35.6659 | -37.187
27.3471 | -35.6661 | -37.192
27.3532 | -35.6664 | -37.1972
27.3593 | -35.6667 | -37.2022
27.3654 | -35.667 | -37.2074
27.3715 | -35.6672 | -37.2125
27.3776 | -35.6675 | -37.2176
27.3837 | -35.6678 | -37.2226
27.3898 | -35.668 | -37.2278
27.3959 | -35.6682 | -37.2329
27.402 | -35.6685 | -37.2383
27.4081 | -35.6687 | -37.2435
27.4142 | -35.6689 | -37.249
27.4203 | -35.6692 | -37.2541
27.4264 | -35.6694 | -37.259
27.4325 | -35.6697 | -37.2642
27.4387 | -35.67 | -37.2692
27.4448 | -35.6702 | -37.2742
27.4509 | -35.6705 | -37.2797
27.457 | -35.6707 | -37.2849
27.4631 | -35.671 | -37.2902
27.4692 | -35.6713 | -37.2955
27.4753 | -35.6715 | -37.3006
27.4814 | -35.6718 | -37.3057
27.4875 | -35.6721 | -37.311
27.4936 | -35.6723 | -37.3162
27.4997 | -35.6726 | -37.3214
27.5058 | -35.6728 | -37.3263
27.5119 | -35.6731 | -37.3315
27.518 | -35.6733 | -37.3366
27.5241 | -35.6736 | -37.3418
27.5302 | -35.6738 | -37.347
27.5363 | -35.6741 | -37.3522
27.5424 | -35.6743 | -37.3573
27.5485 | -35.6746 | -37.3626
27.5546 | -35.6749 | -37.368
27.5607 | -35.6752 | -37.3732
27.5668 | -35.6754 | -37.3785
27.5729 | -35.6757 | -37.3837
27.5791 | -35.6759 | -37.389
27.5852 | -35.6761 | -37.3942
27.5913 | -35.6764 | -37.3991
27.5974 | -35.6767 | -37.4044
27.6035 | -35.6769 | -37.4096
27.6096 | -35.6772 | -37.4147
27.6157 | -35.6774 | -37.4197
27.6218 | -35.6776 | -37.4252
27.6279 | -35.6779 | -37.4303
27.634 | -35.6781 | -37.4355
27.6401 | -35.6784 | -37.4405
27.6462 | -35.6787 | -37.4457
27.6523 | -35.6789 | -37.4504
27.6584 | -35.6791 | -37.4557
27.6645 | -35.6794 | -37.4609
27.6706 | -35.6796 | -37.4662
27.6767 | -35.6799 | -37.4712
27.6828 | -35.6801 | -37.4758
27.6889 | -35.6804 | -37.4808
27.695 | -35.6806 | -37.4857
27.7011 | -35.6808 | -37.4907
27.7072 | -35.6811 | -37.4958
27.7133 | -35.6815 | -37.5011
27.7194 | -35.6817 | -37.506
27.7256 | -35.6819 | -37.5111
27.7317 | -35.6821 | -37.5157
27.7378 | -35.6824 | -37.521
27.7439 | -35.6826 | -37.5263
27.75 | -35.6829 | -37.5314
27.7561 | -35.6831 | -37.5365
27.7622 | -35.6833 | -37.5417
27.7683 | -35.6836 | -37.5472
27.7744 | -35.6838 | -37.5523
27.7805 | -35.6841 | -37.5576
27.7866 | -35.6844 | -37.563
27.7927 | -35.6846 | -37.5679
27.7988 | -35.6849 | -37.5731
27.8049 | -35.6851 | -37.5784
27.811 | -35.6854 | -37.5837
27.8171 | -35.6857 | -37.589
27.8232 | -35.6859 | -37.5938
27.8293 | -35.6862 | -37.599
27.8354 | -35.6864 | -37.6043
27.8415 | -35.6867 | -37.6093
27.8476 | -35.6869 | -37.6148
27.8537 | -35.6872 | -37.6197
27.8598 | -35.6874 | -37.6247
27.866 | -35.6877 | -37.6301
27.8721 | -35.6879 | -37.6352
27.8782 | -35.6882 | -37.6404
27.8843 | -35.6885 | -37.6456
27.8904 | -35.6888 | -37.6505
27.8965 | -35.689 | -37.6554
27.9026 | -35.6893 | -37.6604
27.9087 | -35.6895 | -37.6658
27.9148 | -35.6897 | -37.6708
27.9209 | -35.69 | -37.6759
27.927 | -35.6902 | -37.6809
27.9331 | -35.6905 | -37.686
27.9392 | -35.6908 | -37.6912
27.9453 | -35.691 | -37.6963
27.9514 | -35.6912 | -37.7014
27.9575 | -35.6915 | -37.7069
27.9636 | -35.6918 | -37.712
27.9697 | -35.692 | -37.717
27.9758 | -35.6922 | -37.7222
27.9819 | -35.6925 | -37.7276
27.988 | -35.6928 | -37.7328
27.9941 | -35.693 | -37.738
28.0002 | -35.6933 | -37.7429
28.0063 | -35.6935 | -37.7479
28.0125 | -35.6938 | -37.7529
28.0186 | -35.694 | -37.758
28.0247 | -35.6943 | -37.763
28.0308 | -35.6946 | -37.7678
28.0369 | -35.6948 | -37.7726
28.043 | -35.695 | -37.7781
28.0491 | -35.6953 | -37.7828
28.0552 | -35.6955 | -37.7879
28.0613 | -35.6958 | -37.793
28.0674 | -35.696 | -37.7982
28.0735 | -35.6963 | -37.8031

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.3302 | -42.3954
21.4908 | -35.5171 | -34.9076
16.4948 | -34.8932 | -27.5741
12.6602 | -34.5324 | -21.3255
9.7171 | -34.3169 | -16.4292
7.4582 | -34.1705 | -12.4561
5.7244 | -34.0652 | -9.2391
4.3936 | -34.0031 | -6.6367
3.3722 | -33.987 | -4.5903
2.5883 | -33.9663 | -3.0387
1.9866 | -33.9637 | -1.7981
1.5248 | -33.9764 | -0.80266
1.1703 | -33.9784 | -0.15476
1 | -33.986 | 0.16841
0.89822 | -33.9754 | 0.4381
0.68941 | -33.9598 | 1.0174
0.52919 | -33.9643 | 1.4682
0.40612 | -33.9602 | 2.078
0.31171 | -33.9598 | 2.5296
0.23928 | -33.957 | 3.2624
0.18358 | -33.9778 | 3.6771
0.14091 | -33.9823 | 4.5622
0.10819 | -34.0036 | 5.7198
0.083017 | -34.0324 | 7.1917
0.063705 | -34.0754 | 9.0626
0.048923 | -34.1324 | 11.561
0.037479 | -34.2332 | 14.7494
0.028801 | -34.4235 | 18.749
0.022078 | -34.7217 | 24.063
0.017023 | -35.1986 | 30.4901
0.013018 | -36.0091 | 38.1984
0.010014 | -37.1155 | 46.1075

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003658 | -7.4789
0.007316 | -25.8357
0.010974 | -33.4018
0.014633 | -38.9437
0.018291 | -40.6496
0.021949 | -43.0005
0.025607 | -45.5991
0.029265 | -48.0029
0.032923 | -49.3429
0.036581 | -49.9122
0.04024 | -50.7284
0.043898 | -51.9582
0.047556 | -52.8756
0.051214 | -54.971
0.054872 | -55.5946
0.05853 | -56.5429
0.062188 | -57.685
0.065847 | -57.4189
0.069505 | -58.0254
0.073163 | -58.8558
0.076821 | -60.4003
0.080479 | -60.8752
0.084137 | -62.0886
0.087796 | -61.9501
0.091454 | -61.7403
0.095112 | -61.3106
0.09877 | -62.1292
0.10243 | -63.0087
0.10609 | -62.9398
0.10974 | -64.0238
0.1134 | -63.7105
0.11706 | -65.0889
0.12072 | -66.9586
0.12438 | -67.2261
0.12804 | -67.55
0.13169 | -67.2847
0.13535 | -66.8416
0.13901 | -67.1226
0.14267 | -67.1706
0.14633 | -66.7843
0.14998 | -66.9423
0.15364 | -67.1197
0.1573 | -67.8268
0.16096 | -67.9301
0.16462 | -68.3034
0.16828 | -68.7772
0.17193 | -69.1592
0.17559 | -68.6908
0.17925 | -69.1613
0.18291 | -69.7448
0.18657 | -69.8508
0.19022 | -70.4432
0.19388 | -70.2075
0.19754 | -70.664
0.2012 | -71.0828
0.20486 | -70.9072
0.20851 | -72.6463
0.21217 | -71.2673
0.21583 | -70.9442
0.21949 | -72.2891
0.22315 | -72.7017
0.22681 | -72.9847
0.23046 | -72.8806
0.23412 | -71.5855
0.23778 | -71.8129
0.24144 | -72.2213
0.2451 | -73.36
0.24875 | -74.3797
0.25241 | -74.2198
0.25607 | -74.1555
0.25973 | -73.7201
0.26339 | -73.6891
0.26704 | -74.1308
0.2707 | -75.5375
0.27436 | -75.0042
0.27802 | -73.7685
0.28168 | -74.4173
0.28534 | -74.1857
0.28899 | -74.1151
0.29265 | -74.5581
0.29631 | -75.2752
0.29997 | -75.2114
0.30363 | -76.7882
0.30728 | -75.6999
0.31094 | -75.1093
0.3146 | -75.2352
0.31826 | -75.6237
0.32192 | -75.5498
0.32558 | -75.5651
0.32923 | -75.9943
0.33289 | -75.8981
0.33655 | -75.6859
0.34021 | -75.9259
0.34387 | -76.6889
0.34752 | -77.0835
0.35118 | -76.6251
0.35484 | -76.4047
0.3585 | -76.4321
0.36216 | -77.1123
0.36582 | -76.824
0.36947 | -77.7016
0.37313 | -77.3188
0.37679 | -76.9288
0.38045 | -77.8331
0.3841 | -76.5812
0.38776 | -76.7402
0.39142 | -77.1107
0.39508 | -77.4966
0.39874 | -78.726
0.4024 | -78.8292
0.40605 | -78.2565
0.40971 | -78.6538
0.41337 | -79.1447
0.41703 | -78.0158
0.42069 | -77.7524
0.42435 | -77.5197
0.428 | -79.5038
0.43166 | -79.7661
0.43532 | -79.2193
0.43898 | -79.3421
0.44264 | -79.1062
0.44629 | -78.4298
0.44995 | -78.8845
0.45361 | -78.9379
0.45727 | -78.9918
0.46093 | -79.7424
0.46459 | -80.0572
0.46824 | -79.427
0.4719 | -79.1872
0.47556 | -79.5157
0.47922 | -79.8769
0.48287 | -80.2608
0.48653 | -78.4498
0.49019 | -79.4209
0.49385 | -80.9588
0.49751 | -80.5003
0.50117 | -80.5846
0.50482 | -79.7959
0.50848 | -80.2516
0.51214 | -80.9524
0.5158 | -82.1362
0.51946 | -81.5874
0.52311 | -81.0623
0.52677 | -81.1562
0.53043 | -81.3213
0.53409 | -81.2925
0.53775 | -81.3701
0.54141 | -81.2623
0.54506 | -81.0354
0.54872 | -80.9851
0.55238 | -80.8403
0.55604 | -81.039
0.5597 | -81.4395
0.56336 | -81.7907
0.56701 | -82.3458
0.57067 | -82.1601
0.57433 | -81.5408
0.57799 | -81.5134
0.58164 | -81.8579
0.5853 | -82.4017
0.58896 | -82.7683
0.59262 | -82.5673
0.59628 | -82.3978
0.59994 | -82.4502
0.60359 | -82.8697
0.60725 | -83.216
0.61091 | -82.9244
0.61457 | -82.4954
0.61823 | -82.3368
0.62189 | -82.9378
0.62554 | -83.431
0.6292 | -82.9587
0.63286 | -82.6779
0.63652 | -83.0357
0.64018 | -83.3523
0.64383 | -83.1685
0.64749 | -82.947
0.65115 | -82.8746
0.65481 | -82.8381
0.65847 | -82.7588
0.66212 | -82.8219
0.66578 | -83.1038
0.66944 | -83.5081
0.6731 | -84.1889
0.67676 | -84.6942
0.68041 | -85.0002
0.68407 | -84.8231
0.68773 | -83.8604
0.69139 | -83.0067
0.69505 | -83.0336
0.69871 | -83.566
0.70236 | -84.1083
0.70602 | -84.3714
0.70968 | -84.1212
0.71334 | -83.8912
0.717 | -84.0622
0.72066 | -84.3928
0.72431 | -84.1882
0.72797 | -83.4854
0.73163 | -83.2788
0.73529 | -83.734
0.73895 | -84.5674
0.7426 | -84.9881
0.74626 | -84.6738
0.74992 | -84.6442
0.75358 | -84.3412
0.75724 | -83.9192
0.76089 | -84.2227
0.76455 | -84.6745
0.76821 | -84.6341
0.77187 | -84.8006
0.77553 | -85.0354
0.77919 | -84.5842
0.78284 | -84.4002
0.7865 | -84.8915
0.79016 | -85.0501
0.79382 | -84.7193
0.79748 | -84.263
0.80113 | -84.136
0.80479 | -84.7911
0.80845 | -85.2212
0.81211 | -85.193
0.81577 | -85.0436
0.81942 | -84.8929
0.82308 | -85.1069
0.82674 | -85.3338
0.8304 | -85.4912
0.83406 | -85.7693
0.83771 | -85.7606
0.84137 | -85.2207
0.84503 | -85.2427
0.84869 | -85.4394
0.85235 | -85.2856
0.85601 | -85.1855
0.85966 | -85.1297
0.86332 | -85.446
0.86698 | -85.8571
0.87064 | -85.9711
0.8743 | -86.2719
0.87796 | -86.3312
0.88161 | -85.9549
0.88527 | -85.3499
0.88893 | -84.8482
0.89259 | -84.7938
0.89625 | -84.7998
0.8999 | -84.6194
0.90356 | -84.5688
0.90722 | -85.0296
0.91088 | -85.809
0.91454 | -86.3082
0.91819 | -86.4344
0.92185 | -86.354
0.92551 | -86.459
0.92917 | -86.4441
0.93283 | -86.3258
0.93649 | -86.4558
0.94014 | -86.6847
0.9438 | -86.4396
0.94746 | -86.1993
0.95112 | -86.0002
0.95478 | -85.7594
0.95843 | -85.9571
0.96209 | -86.2037
0.96575 | -85.9822
0.96941 | -86.0602
0.97307 | -86.7385
0.97672 | -86.7775
0.98038 | -86.4951
0.98404 | -86.7988
0.9877 | -87.0274
0.99136 | -86.9625
0.99502 | -86.7752
0.99867 | -86.4748
1.0023 | -86.513
1.006 | -86.6545
1.0096 | -86.5919
1.0133 | -86.5111
1.017 | -86.5905
1.0206 | -86.4239
1.0243 | -86.6901
1.0279 | -87.6018
1.0316 | -88.0672
1.0353 | -87.664
1.0389 | -87.2382
1.0426 | -87.2658
1.0462 | -87.6537
1.0499 | -88.0828
1.0535 | -87.7735
1.0572 | -87.5047
1.0609 | -87.3685
1.0645 | -87.5745
1.0682 | -88.0468
1.0718 | -88.4569
1.0755 | -88.61
1.0792 | -88.5602
1.0828 | -88.3186
1.0865 | -87.9107
1.0901 | -87.482
1.0938 | -87.1994
1.0974 | -87.1595
1.1011 | -87.3398
1.1048 | -87.6344
1.1084 | -87.9261
1.1121 | -88.0676
1.1157 | -87.9727
1.1194 | -87.7113
1.1231 | -87.6449
1.1267 | -88.0218
1.1304 | -88.5384
1.134 | -88.7806
1.1377 | -88.7142
1.1413 | -88.4113
1.145 | -88.1204
1.1487 | -88.1469
1.1523 | -88.4041
1.156 | -88.5072
1.1596 | -88.4252
1.1633 | -88.4096
1.1669 | -88.447
1.1706 | -88.4138
1.1743 | -88.4107
1.1779 | -88.5138
1.1816 | -88.6124
1.1852 | -88.6615
1.1889 | -88.7209
1.1926 | -88.8158
1.1962 | -88.8593
1.1999 | -88.7924
1.2035 | -88.7328
1.2072 | -88.8137
1.2108 | -89.0486
1.2145 | -89.2509
1.2182 | -89.2626
1.2218 | -89.2526
1.2255 | -89.2466
1.2291 | -89.1533
1.2328 | -89.1134
1.2365 | -89.1883
1.2401 | -89.1942
1.2438 | -89.0862
1.2474 | -89.1293
1.2511 | -89.4132
1.2547 | -89.526
1.2584 | -89.2001
1.2621 | -88.8014
1.2657 | -88.7377
1.2694 | -89.0002
1.273 | -89.2143
1.2767 | -89.1039
1.2804 | -88.8936
1.284 | -88.9511
1.2877 | -89.2921
1.2913 | -89.6221
1.295 | -89.7633
1.2986 | -89.7901
1.3023 | -89.813
1.306 | -89.8132
1.3096 | -89.7196
1.3133 | -89.5644
1.3169 | -89.4368
1.3206 | -89.4137
1.3242 | -89.5611
1.3279 | -89.9547
1.3316 | -90.5266
1.3352 | -90.9296
1.3389 | -90.8421
1.3425 | -90.4108
1.3462 | -89.9898
1.3499 | -89.7295
1.3535 | -89.6341
1.3572 | -89.7521
1.3608 | -90.0568
1.3645 | -90.2597
1.3681 | -90.1217
1.3718 | -89.9473
1.3755 | -90.1225
1.3791 | -90.629
1.3828 | -91.0051
1.3864 | -90.7451
1.3901 | -90.0072
1.3938 | -89.4447
1.3974 | -89.4122
1.4011 | -89.7357
1.4047 | -90.0336
1.4084 | -90.2127
1.412 | -90.466
1.4157 | -90.8613
1.4194 | -91.18
1.423 | -91.321
1.4267 | -91.3939
1.4303 | -91.3102
1.434 | -91.0413
1.4377 | -90.7664
1.4413 | -90.5797
1.445 | -90.5043
1.4486 | -90.6035
1.4523 | -90.8405
1.4559 | -90.9575
1.4596 | -90.8201
1.4633 | -90.7611
1.4669 | -90.9494
1.4706 | -91.0976
1.4742 | -91.0879
1.4779 | -91.1378
1.4815 | -91.2771
1.4852 | -91.3435
1.4889 | -91.2425
1.4925 | -91.0106
1.4962 | -90.8047
1.4998 | -90.7276
1.5035 | -90.69
1.5072 | -90.6376
1.5108 | -90.7005
1.5145 | -91.0632
1.5181 | -91.7156
1.5218 | -92.3031
1.5254 | -92.4632
1.5291 | -92.1787
1.5328 | -91.7139
1.5364 | -91.4275
1.5401 | -91.4387
1.5437 | -91.6342
1.5474 | -91.8622
1.5511 | -91.9675
1.5547 | -91.9656
1.5584 | -92.0153
1.562 | -92.0406
1.5657 | -91.7973
1.5693 | -91.3545
1.573 | -91.1078
1.5767 | -91.2459
1.5803 | -91.5559
1.584 | -91.7974
1.5876 | -92.0165
1.5913 | -92.1935
1.595 | -92.2859
1.5986 | -92.2836
1.6023 | -92.2429
1.6059 | -92.2852
1.6096 | -92.4591
1.6132 | -92.6836
1.6169 | -92.8023
1.6206 | -92.7251
1.6242 | -92.5336
1.6279 | -92.3197
1.6315 | -92.1287
1.6352 | -91.9903
1.6388 | -91.8823
1.6425 | -91.7233
1.6462 | -91.5006
1.6498 | -91.3284
1.6535 | -91.2971
1.6571 | -91.4069
1.6608 | -91.5918
1.6645 | -91.7558
1.6681 | -91.801
1.6718 | -91.7671
1.6754 | -91.8022
1.6791 | -91.9489
1.6827 | -92.0845
1.6864 | -92.1034
1.6901 | -91.9869
1.6937 | -91.8269
1.6974 | -91.7054
1.701 | -91.6901
1.7047 | -91.7492
1.7084 | -91.8321
1.712 | -91.915
1.7157 | -91.9985
1.7193 | -92.0568
1.723 | -92.0672
1.7266 | -92.0015
1.7303 | -91.9321
1.734 | -91.9341
1.7376 | -92.0315
1.7413 | -92.1521
1.7449 | -92.2076
1.7486 | -92.1377
1.7523 | -91.9565
1.7559 | -91.725
1.7596 | -91.5334
1.7632 | -91.4561
1.7669 | -91.5024
1.7705 | -91.6368
1.7742 | -91.8505
1.7779 | -92.1568
1.7815 | -92.4648
1.7852 | -92.5934
1.7888 | -92.5095
1.7925 | -92.3126
1.7961 | -92.1063
1.7998 | -91.9337
1.8035 | -91.7928
1.8071 | -91.6908
1.8108 | -91.582
1.8144 | -91.4486
1.8181 | -91.3342
1.8218 | -91.3409
1.8254 | -91.4949
1.8291 | -91.6711
1.8327 | -91.7542
1.8364 | -91.7888
1.84 | -91.8885
1.8437 | -92.1061
1.8474 | -92.3621
1.851 | -92.5514
1.8547 | -92.6127
1.8583 | -92.6282
1.862 | -92.7032
1.8657 | -92.81
1.8693 | -92.8331
1.873 | -92.7327
1.8766 | -92.5897
1.8803 | -92.469
1.8839 | -92.3794
1.8876 | -92.307
1.8913 | -92.2519
1.8949 | -92.2603
1.8986 | -92.381
1.9022 | -92.5851
1.9059 | -92.785
1.9096 | -92.9034
1.9132 | -92.9431
1.9169 | -92.9398
1.9205 | -92.8924
1.9242 | -92.8231
1.9278 | -92.7902
1.9315 | -92.8498
1.9352 | -92.9858
1.9388 | -93.1121
1.9425 | -93.1852
1.9461 | -93.2531
1.9498 | -93.3288
1.9534 | -93.3818
1.9571 | -93.3568
1.9608 | -93.264
1.9644 | -93.1478
1.9681 | -93.0416
1.9717 | -92.9886
1.9754 | -92.9929
1.9791 | -93.0454
1.9827 | -93.1715
1.9864 | -93.3363
1.99 | -93.4879
1.9937 | -93.5888
1.9973 | -93.6135
2.001 | -93.5307
2.0047 | -93.3185
2.0083 | -93.1146
2.012 | -93.0582
2.0156 | -93.0971
2.0193 | -93.0831
2.023 | -92.9663
2.0266 | -92.8818
2.0303 | -92.9362
2.0339 | -93.1303
2.0376 | -93.3904
2.0412 | -93.5822
2.0449 | -93.635
2.0486 | -93.6092
2.0522 | -93.5995
2.0559 | -93.6198
2.0595 | -93.6131
2.0632 | -93.5676
2.0669 | -93.5075
2.0705 | -93.4305
2.0742 | -93.3234
2.0778 | -93.2039
2.0815 | -93.1518
2.0851 | -93.1912
2.0888 | -93.2893
2.0925 | -93.3985
2.0961 | -93.5089
2.0998 | -93.6433
2.1034 | -93.8255
2.1071 | -94.0422
2.1107 | -94.2179
2.1144 | -94.3221
2.1181 | -94.3049
2.1217 | -94.1606
2.1254 | -93.9183
2.129 | -93.6462
2.1327 | -93.4236
2.1364 | -93.2925
2.14 | -93.2703
2.1437 | -93.3199
2.1473 | -93.3718
2.151 | -93.3896
2.1546 | -93.3538
2.1583 | -93.26
2.162 | -93.1853
2.1656 | -93.191
2.1693 | -93.3078
2.1729 | -93.4855
2.1766 | -93.6294
2.1803 | -93.6508
2.1839 | -93.5431
2.1876 | -93.379
2.1912 | -93.2637
2.1949 | -93.2689
2.1985 | -93.3962
2.2022 | -93.5837
2.2059 | -93.7215
2.2095 | -93.735
2.2132 | -93.6628
2.2168 | -93.6025
2.2205 | -93.6229
2.2242 | -93.7522
2.2278 | -93.9343
2.2315 | -94.0889
2.2351 | -94.1568
2.2388 | -94.1446
2.2424 | -94.0709
2.2461 | -93.9933
2.2498 | -93.9311
2.2534 | -93.9061
2.2571 | -93.9003
2.2607 | -93.8978
2.2644 | -93.8603
2.2681 | -93.7913
2.2717 | -93.7214
2.2754 | -93.697
2.279 | -93.734
2.2827 | -93.8387
2.2863 | -93.9895
2.29 | -94.1264
2.2937 | -94.219
2.2973 | -94.251
2.301 | -94.2072
2.3046 | -94.106
2.3083 | -94.0123
2.3119 | -93.9814
2.3156 | -94.0334
2.3193 | -94.1221
2.3229 | -94.1602
2.3266 | -94.0953
2.3302 | -94.0003
2.3339 | -93.9552
2.3376 | -94.02
2.3412 | -94.1713
2.3449 | -94.3416
2.3485 | -94.4552
2.3522 | -94.5114
2.3558 | -94.5312
2.3595 | -94.541
2.3632 | -94.5383
2.3668 | -94.5263
2.3705 | -94.5115
2.3741 | -94.5401
2.3778 | -94.6324
2.3815 | -94.7568
2.3851 | -94.8506
2.3888 | -94.8253
2.3924 | -94.6627
2.3961 | -94.4358
2.3997 | -94.2439
2.4034 | -94.1607
2.4071 | -94.1936
2.4107 | -94.3006
2.4144 | -94.4353
2.418 | -94.5389
2.4217 | -94.5704
2.4254 | -94.5053
2.429 | -94.3584
2.4327 | -94.1639
2.4363 | -93.9769
2.44 | -93.8422
2.4436 | -93.805
2.4473 | -93.8774
2.451 | -94.0491
2.4546 | -94.2915
2.4583 | -94.5792
2.4619 | -94.8917
2.4656 | -95.2063
2.4692 | -95.4917
2.4729 | -95.6989
2.4766 | -95.7832
2.4802 | -95.679
2.4839 | -95.3672
2.4875 | -94.9199
2.4912 | -94.4697
2.4949 | -94.1419
2.4985 | -93.989
2.5022 | -93.9877
2.5058 | -94.0931
2.5095 | -94.2733
2.5131 | -94.5057
2.5168 | -94.7586
2.5205 | -94.9818
2.5241 | -95.1398
2.5278 | -95.2264
2.5314 | -95.2398
2.5351 | -95.1849
2.5388 | -95.0911
2.5424 | -95.0134
2.5461 | -94.984
2.5497 | -95.0043
2.5534 | -95.0505
2.557 | -95.0762
2.5607 | -95.0765
2.5644 | -95.075
2.568 | -95.1352
2.5717 | -95.285
2.5753 | -95.5039
2.579 | -95.7011
2.5827 | -95.7823
2.5863 | -95.7554
2.59 | -95.6707
2.5936 | -95.5601
2.5973 | -95.445
2.6009 | -95.3336
2.6046 | -95.2463
2.6083 | -95.2164
2.6119 | -95.2558
2.6156 | -95.3375
2.6192 | -95.4323
2.6229 | -95.4987
2.6265 | -95.5248
2.6302 | -95.5342
2.6339 | -95.5546
2.6375 | -95.5976
2.6412 | -95.6418
2.6448 | -95.6709
2.6485 | -95.6799
2.6522 | -95.6608
2.6558 | -95.5911
2.6595 | -95.4537
2.6631 | -95.2807
2.6668 | -95.1212
2.6704 | -95.0004
2.6741 | -94.9359
2.6778 | -94.9249
2.6814 | -94.9499
2.6851 | -94.9765
2.6887 | -94.994
2.6924 | -94.9965
2.6961 | -94.9856
2.6997 | -94.9628
2.7034 | -94.9419
2.707 | -94.9384
2.7107 | -94.9386
2.7143 | -94.9503
2.718 | -95.003
2.7217 | -95.1115
2.7253 | -95.28
2.729 | -95.4847
2.7326 | -95.6805
2.7363 | -95.8135
2.74 | -95.8458
2.7436 | -95.7604
2.7473 | -95.6132
2.7509 | -95.4623
2.7546 | -95.3306
2.7582 | -95.2447
2.7619 | -95.2196
2.7656 | -95.2625
2.7692 | -95.3658
2.7729 | -95.5131
2.7765 | -95.6722
2.7802 | -95.8098
2.7838 | -95.9045
2.7875 | -95.9571
2.7912 | -95.9993
2.7948 | -96.0443
2.7985 | -96.1161
2.8021 | -96.2186
2.8058 | -96.327
2.8095 | -96.4012
2.8131 | -96.4221
2.8168 | -96.3701
2.8204 | -96.2575
2.8241 | -96.1193
2.8277 | -95.9856
2.8314 | -95.876
2.8351 | -95.7922
2.8387 | -95.7119
2.8424 | -95.6365
2.846 | -95.5903
2.8497 | -95.5962
2.8534 | -95.6372
2.857 | -95.7026
2.8607 | -95.7678
2.8643 | -95.8252
2.868 | -95.8742
2.8716 | -95.9109
2.8753 | -95.9253
2.879 | -95.925
2.8826 | -95.9279
2.8863 | -95.9497
2.8899 | -95.9819
2.8936 | -96.0125
2.8973 | -96.0529
2.9009 | -96.1196
2.9046 | -96.1935
2.9082 | -96.2619
2.9119 | -96.3095
2.9155 | -96.3125
2.9192 | -96.2999
2.9229 | -96.288
2.9265 | -96.2842
2.9302 | -96.2897
2.9338 | -96.2981
2.9375 | -96.2937
2.9411 | -96.2635
2.9448 | -96.1881
2.9485 | -96.0924
2.9521 | -96.0121
2.9558 | -95.9767
2.9594 | -96.0138
2.9631 | -96.1275
2.9668 | -96.2986
2.9704 | -96.482
2.9741 | -96.6301
2.9777 | -96.7237
2.9814 | -96.7569
2.985 | -96.7463
2.9887 | -96.7097
2.9924 | -96.6682
2.996 | -96.6253
2.9997 | -96.5976
3.0033 | -96.598
3.007 | -96.6035
3.0107 | -96.5928
3.0143 | -96.5353
3.018 | -96.4404
3.0216 | -96.3075
3.0253 | -96.1629
3.0289 | -96.0436
3.0326 | -95.9594
3.0363 | -95.9209
3.0399 | -95.9551
3.0436 | -96.0759
3.0472 | -96.2789
3.0509 | -96.523
3.0546 | -96.7348
3.0582 | -96.8477
3.0619 | -96.8427
3.0655 | -96.7569
3.0692 | -96.6665
3.0728 | -96.6211
3.0765 | -96.644
3.0802 | -96.7474
3.0838 | -96.9041
3.0875 | -97.0674
3.0911 | -97.1844
3.0948 | -97.2112
3.0985 | -97.1386
3.1021 | -96.9861
3.1058 | -96.8064
3.1094 | -96.6496
3.1131 | -96.5289
3.1167 | -96.434
3.1204 | -96.3453
3.1241 | -96.2783
3.1277 | -96.2416
3.1314 | -96.2502
3.135 | -96.3038
3.1387 | -96.3893
3.1423 | -96.4795
3.146 | -96.5616
3.1497 | -96.6196
3.1533 | -96.6316
3.157 | -96.6285
3.1606 | -96.6271
3.1643 | -96.6513
3.168 | -96.6694
3.1716 | -96.7074
3.1753 | -96.7459
3.1789 | -96.7825
3.1826 | -96.8074
3.1862 | -96.8187
3.1899 | -96.7986
3.1936 | -96.7527
3.1972 | -96.6833
3.2009 | -96.6175
3.2045 | -96.5696
3.2082 | -96.5398
3.2119 | -96.563
3.2155 | -96.6332
3.2192 | -96.727
3.2228 | -96.8342
3.2265 | -96.9236
3.2301 | -96.9912
3.2338 | -97.0422
3.2375 | -97.0716
3.2411 | -97.0774
3.2448 | -97.0608
3.2484 | -97.0126
3.2521 | -96.9434
3.2557 | -96.8439
3.2594 | -96.7322
3.2631 | -96.6303
3.2667 | -96.5534
3.2704 | -96.5329
3.274 | -96.5714
3.2777 | -96.6543
3.2814 | -96.7658
3.285 | -96.8885
3.2887 | -97.011
3.2923 | -97.1134
3.296 | -97.1824
3.2996 | -97.2071
3.3033 | -97.171
3.307 | -97.0935
3.3106 | -97.024
3.3143 | -96.9768
3.3179 | -96.9568
3.3216 | -96.957
3.3253 | -96.9773
3.3289 | -97.0058
3.3326 | -97.0352
3.3362 | -97.0564
3.3399 | -97.0674
3.3435 | -97.0728
3.3472 | -97.07
3.3509 | -97.049
3.3545 | -97.0242
3.3582 | -97.0049
3.3618 | -96.9865
3.3655 | -96.968
3.3692 | -96.9525
3.3728 | -96.9443
3.3765 | -96.9244
3.3801 | -96.8788
3.3838 | -96.8163
3.3874 | -96.7385
3.3911 | -96.675
3.3948 | -96.6434
3.3984 | -96.6395
3.4021 | -96.6681
3.4057 | -96.7237
3.4094 | -96.781
3.4131 | -96.8333
3.4167 | -96.8712
3.4204 | -96.9064
3.424 | -96.9439
3.4277 | -96.9755
3.4313 | -97.0155
3.435 | -97.0471
3.4387 | -97.0816
3.4423 | -97.1379
3.446 | -97.2095
3.4496 | -97.3023
3.4533 | -97.4259
3.4569 | -97.5552
3.4606 | -97.6509
3.4643 | -97.711
3.4679 | -97.7411
3.4716 | -97.7526
3.4752 | -97.7545
3.4789 | -97.7515
3.4826 | -97.746
3.4862 | -97.7274
3.4899 | -97.7168
3.4935 | -97.7128
3.4972 | -97.7202
3.5008 | -97.744
3.5045 | -97.7793
3.5082 | -97.8244
3.5118 | -97.8675
3.5155 | -97.9033
3.5191 | -97.9216
3.5228 | -97.9143
3.5265 | -97.8877
3.5301 | -97.8373
3.5338 | -97.7951
3.5374 | -97.7717
3.5411 | -97.7788
3.5447 | -97.8232
3.5484 | -97.903
3.5521 | -98.0152
3.5557 | -98.13
3.5594 | -98.2182
3.563 | -98.2494
3.5667 | -98.2273
3.5704 | -98.1667
3.574 | -98.09
3.5777 | -98.0275
3.5813 | -97.9887
3.585 | -97.9968
3.5886 | -98.047
3.5923 | -98.1421
3.596 | -98.2718
3.5996 | -98.4191
3.6033 | -98.5455
3.6069 | -98.6118
3.6106 | -98.6048
3.6142 | -98.5238
3.6179 | -98.4
3.6216 | -98.2613
3.6252 | -98.1449
3.6289 | -98.0723
3.6325 | -98.0495
3.6362 | -98.0717
3.6399 | -98.1327
3.6435 | -98.2187
3.6472 | -98.3135
3.6508 | -98.3939
3.6545 | -98.4448
3.6581 | -98.4637
3.6618 | -98.4512
3.6655 | -98.4119
3.6691 | -98.3592
3.6728 | -98.3022
3.6764 | -98.2451
3.6801 | -98.1825
3.6838 | -98.1287
3.6874 | -98.0966
3.6911 | -98.1265
3.6947 | -98.1477
3.6984 | -98.1803
3.702 | -98.2206
3.7057 | -98.2558
3.7094 | -98.275
3.713 | -98.2637
3.7167 | -98.2248
3.7203 | -98.1673
3.724 | -98.1006
3.7277 | -98.0406
3.7313 | -98.0132
3.735 | -98.0234
3.7386 | -98.0689
3.7423 | -98.1458
3.7459 | -98.2391
3.7496 | -98.3452
3.7533 | -98.4419
3.7569 | -98.5131
3.7606 | -98.5501
3.7642 | -98.5476
3.7679 | -98.499
3.7715 | -98.4228
3.7752 | -98.3314
3.7789 | -98.2473
3.7825 | -98.1798
3.7862 | -98.1474
3.7898 | -98.1476
3.7935 | -98.1731
3.7972 | -98.2251
3.8008 | -98.298
3.8045 | -98.3681
3.8081 | -98.4393
3.8118 | -98.4901
3.8154 | -98.5148
3.8191 | -98.5118
3.8228 | -98.4897
3.8264 | -98.4569
3.8301 | -98.4222
3.8337 | -98.3929
3.8374 | -98.3815
3.8411 | -98.3744
3.8447 | -98.3648
3.8484 | -98.329
3.852 | -98.2742
3.8557 | -98.2061
3.8593 | -98.1285
3.863 | -98.0477
3.8667 | -97.9785
3.8703 | -97.9219
3.874 | -97.888
3.8776 | -97.8811
3.8813 | -97.9056
3.885 | -97.9643
3.8886 | -98.0666
3.8923 | -98.1992
3.8959 | -98.3396
3.8996 | -98.4674
3.9032 | -98.5658
3.9069 | -98.6184
3.9106 | -98.6314
3.9142 | -98.5901
3.9179 | -98.5274
3.9215 | -98.4612
3.9252 | -98.4077
3.9288 | -98.3782
3.9325 | -98.3802
3.9362 | -98.4312
3.9398 | -98.5235
3.9435 | -98.6356
3.9471 | -98.7536
3.9508 | -98.8566
3.9545 | -98.9208
3.9581 | -98.9289
3.9618 | -98.8756
3.9654 | -98.7831
3.9691 | -98.6683
3.9727 | -98.5561
3.9764 | -98.4702
3.9801 | -98.4009
3.9837 | -98.3613
3.9874 | -98.3551
3.991 | -98.3834
3.9947 | -98.4315
3.9984 | -98.4923
4.002 | -98.5609
4.0057 | -98.6235
4.0093 | -98.6643
4.013 | -98.6809
4.0166 | -98.6718
4.0203 | -98.6427
4.024 | -98.6048
4.0276 | -98.5405
4.0313 | -98.4559
4.0349 | -98.3542
4.0386 | -98.2558
4.0423 | -98.1815
4.0459 | -98.149
4.0496 | -98.1602
4.0532 | -98.2271
4.0569 | -98.3521
4.0605 | -98.508
4.0642 | -98.6804
4.0679 | -98.8464
4.0715 | -98.9718
4.0752 | -99.0353
4.0788 | -99.0442
4.0825 | -98.9969
4.0861 | -98.9167
4.0898 | -98.849
4.0935 | -98.8084
4.0971 | -98.7923
4.1008 | -98.7977
4.1044 | -98.8309
4.1081 | -98.8723
4.1118 | -98.9268
4.1154 | -98.9624
4.1191 | -98.984
4.1227 | -98.9883
4.1264 | -98.967
4.13 | -98.9166
4.1337 | -98.8634
4.1374 | -98.8236
4.141 | -98.8021
4.1447 | -98.8093
4.1483 | -98.8396
4.152 | -98.8801
4.1557 | -98.93
4.1593 | -98.9944
4.163 | -99.0587
4.1666 | -99.1065
4.1703 | -99.1455
4.1739 | -99.171
4.1776 | -99.1922
4.1813 | -99.2076
4.1849 | -99.2239
4.1886 | -99.2471
4.1922 | -99.2869
4.1959 | -99.3286
4.1996 | -99.3655
4.2032 | -99.3847
4.2069 | -99.3801
4.2105 | -99.3587
4.2142 | -99.3277
4.2178 | -99.2885
4.2215 | -99.2695
4.2252 | -99.2368
4.2288 | -99.2147
4.2325 | -99.2052
4.2361 | -99.1908
4.2398 | -99.1657
4.2435 | -99.1506
4.2471 | -99.1418
4.2508 | -99.1504
4.2544 | -99.1762
4.2581 | -99.2077
4.2617 | -99.2414
4.2654 | -99.2756
4.2691 | -99.304
4.2727 | -99.3193
4.2764 | -99.3151
4.28 | -99.2975
4.2837 | -99.2679
4.2873 | -99.23
4.291 | -99.1945
4.2947 | -99.174
4.2983 | -99.1558
4.302 | -99.1548
4.3056 | -99.1692
4.3093 | -99.1991
4.313 | -99.2247
4.3166 | -99.2377
4.3203 | -99.2308
4.3239 | -99.2137
4.3276 | -99.1962
4.3312 | -99.183
4.3349 | -99.181
4.3386 | -99.1979
4.3422 | -99.2366
4.3459 | -99.2839
4.3495 | -99.3398
4.3532 | -99.3997
4.3569 | -99.4507
4.3605 | -99.4857
4.3642 | -99.5058
4.3678 | -99.5162
4.3715 | -99.5231
4.3751 | -99.5346
4.3788 | -99.5556
4.3825 | -99.5861
4.3861 | -99.6241
4.3898 | -99.6603
4.3934 | -99.6883
4.3971 | -99.7058
4.4008 | -99.7075
4.4044 | -99.6953
4.4081 | -99.6684
4.4117 | -99.6276
4.4154 | -99.5895
4.419 | -99.5513
4.4227 | -99.5292
4.4264 | -99.524
4.43 | -99.5302
4.4337 | -99.5559
4.4373 | -99.5836
4.441 | -99.5896
4.4446 | -99.5791
4.4483 | -99.5432
4.452 | -99.4914
4.4556 | -99.4342
4.4593 | -99.3669
4.4629 | -99.2914
4.4666 | -99.223
4.4703 | -99.1776
4.4739 | -99.1515
4.4776 | -99.1535
4.4812 | -99.1845
4.4849 | -99.2399
4.4885 | -99.3199
4.4922 | -99.4163
4.4959 | -99.5224
4.4995 | -99.6403
4.5032 | -99.7499
4.5068 | -99.8546
4.5105 | -99.9473
4.5142 | -100.0096
4.5178 | -100.0428
4.5215 | -100.0433
4.5251 | -100.0006
4.5288 | -99.9331
4.5324 | -99.8569
4.5361 | -99.7761
4.5398 | -99.6902
4.5434 | -99.6196
4.5471 | -99.5636
4.5507 | -99.5175
4.5544 | -99.4754
4.558 | -99.4325
4.5617 | -99.3957
4.5654 | -99.3665
4.569 | -99.3318
4.5727 | -99.3104
4.5763 | -99.3108
4.58 | -99.3382
4.5837 | -99.3903
4.5873 | -99.4525
4.591 | -99.5317
4.5946 | -99.6273
4.5983 | -99.7132
4.6019 | -99.7814
4.6056 | -99.8331
4.6093 | -99.872
4.6129 | -99.8867
4.6166 | -99.8864
4.6202 | -99.8913
4.6239 | -99.9083
4.6276 | -99.9338
4.6312 | -99.9694
4.6349 | -100.0176
4.6385 | -100.0733
4.6422 | -100.1116
4.6458 | -100.132
4.6495 | -100.1286
4.6532 | -100.1023
4.6568 | -100.0699
4.6605 | -100.023
4.6641 | -99.9709
4.6678 | -99.9379
4.6715 | -99.9208
4.6751 | -99.9281
4.6788 | -99.9556
4.6824 | -100.0132
4.6861 | -100.0945
4.6897 | -100.1785
4.6934 | -100.2489
4.6971 | -100.2923
4.7007 | -100.3041
4.7044 | -100.2888
4.708 | -100.2528
4.7117 | -100.205
4.7153 | -100.1501
4.719 | -100.1109
4.7227 | -100.0888
4.7263 | -100.0744
4.73 | -100.0763
4.7336 | -100.0838
4.7373 | -100.0911
4.741 | -100.0888
4.7446 | -100.0663
4.7483 | -99.9779
4.7519 | -99.9236
4.7556 | -99.8602
4.7592 | -99.7922
4.7629 | -99.739
4.7666 | -99.7226
4.7702 | -99.7465
4.7739 | -99.8048
4.7775 | -99.8977
4.7812 | -100.027
4.7849 | -100.1801
4.7885 | -100.34
4.7922 | -100.4944
4.7958 | -100.6389
4.7995 | -100.7694
4.8031 | -100.8769
4.8068 | -100.9563
4.8105 | -100.9968
4.8141 | -101.01
4.8178 | -100.9991
4.8214 | -100.9596
4.8251 | -100.9008
4.8288 | -100.8322
4.8324 | -100.7532
4.8361 | -100.6703
4.8397 | -100.5773
4.8434 | -100.4823
4.847 | -100.391
4.8507 | -100.3025
4.8544 | -100.217
4.858 | -100.1434
4.8617 | -100.0813
4.8653 | -100.0362
4.869 | -100.0059
4.8727 | -99.9909
4.8763 | -99.9844
4.88 | -99.9881
4.8836 | -99.9976
4.8873 | -100.0077
4.8909 | -100.0211
4.8946 | -100.0374
4.8983 | -100.0627
4.9019 | -100.0897
4.9056 | -100.1169
4.9092 | -100.15
4.9129 | -100.1924
4.9165 | -100.2441
4.9202 | -100.2931
4.9239 | -100.3419
4.9275 | -100.4014
4.9312 | -100.4603
4.9348 | -100.5086
4.9385 | -100.5553
4.9422 | -100.5959
4.9458 | -100.6381
4.9495 | -100.6804
4.9531 | -100.7264
4.9568 | -100.7719
4.9604 | -100.8084
4.9641 | -100.8187
4.9678 | -100.7953
4.9714 | -100.7364
4.9751 | -100.6528
4.9787 | -100.5468
4.9824 | -100.4259
4.9861 | -100.2948
4.9897 | -100.1732
4.9934 | -100.064
4.997 | -99.9659
5.0007 | -99.8911
5.0043 | -99.8355
5.008 | -99.8056
5.0117 | -99.801
5.0153 | -99.8182
5.019 | -99.8485
5.0226 | -99.8911
5.0263 | -99.9499
5.03 | -100.0131
5.0336 | -100.0751
5.0373 | -100.1453
5.0409 | -100.2136
5.0446 | -100.282
5.0482 | -100.3502
5.0519 | -100.4162
5.0556 | -100.4871
5.0592 | -100.5524
5.0629 | -100.6221
5.0665 | -100.6815
5.0702 | -100.7415
5.0738 | -100.8014
5.0775 | -100.8468
5.0812 | -100.8748
5.0848 | -100.8942
5.0885 | -100.898
5.0921 | -100.8951
5.0958 | -100.8931
5.0995 | -100.885
5.1031 | -100.8716
5.1068 | -100.864
5.1104 | -100.8539
5.1141 | -100.829
5.1177 | -100.8077
5.1214 | -100.7799
5.1251 | -100.751
5.1287 | -100.7247
5.1324 | -100.7051
5.136 | -100.6956
5.1397 | -100.689
5.1434 | -100.6827
5.147 | -100.6751
5.1507 | -100.6706
5.1543 | -100.666
5.158 | -100.6634
5.1616 | -100.66
5.1653 | -100.6627
5.169 | -100.6746
5.1726 | -100.6987
5.1763 | -100.7381
5.1799 | -100.7845
5.1836 | -100.8233
5.1873 | -100.8564
5.1909 | -100.8859
5.1946 | -100.8885
5.1982 | -100.8767
5.2019 | -100.8514
5.2055 | -100.8013
5.2092 | -100.7411
5.2129 | -100.674
5.2165 | -100.6088
5.2202 | -100.5483
5.2238 | -100.496
5.2275 | -100.4553
5.2311 | -100.4265
5.2348 | -100.4151
5.2385 | -100.4277
5.2421 | -100.4614
5.2458 | -100.5094
5.2494 | -100.5721
5.2531 | -100.6428
5.2568 | -100.7143
5.2604 | -100.7866
5.2641 | -100.8474
5.2677 | -100.8913
5.2714 | -100.9224
5.275 | -100.9236
5.2787 | -100.9368
5.2824 | -100.9489
5.286 | -100.9625
5.2897 | -100.9711
5.2933 | -100.9832
5.297 | -100.9956
5.3007 | -101.0133
5.3043 | -101.029
5.308 | -101.0412
5.3116 | -101.056
5.3153 | -101.0677
5.3189 | -101.0685
5.3226 | -101.0734
5.3263 | -101.0801
5.3299 | -101.0858
5.3336 | -101.1018
5.3372 | -101.131
5.3409 | -101.157
5.3446 | -101.1841
5.3482 | -101.2173
5.3519 | -101.2494
5.3555 | -101.2735
5.3592 | -101.2936
5.3628 | -101.3072
5.3665 | -101.3146
5.3702 | -101.3222
5.3738 | -101.3255
5.3775 | -101.3283
5.3811 | -101.3338
5.3848 | -101.3472
5.3884 | -101.3662
5.3921 | -101.3935
5.3958 | -101.4263
5.3994 | -101.4609
5.4031 | -101.5041
5.4067 | -101.5449
5.4104 | -101.5781
5.4141 | -101.6
5.4177 | -101.6008
5.4214 | -101.5799
5.425 | -101.5443
5.4287 | -101.4863
5.4323 | -101.411
5.436 | -101.3312
5.4397 | -101.2535
5.4433 | -101.1876
5.447 | -101.138
5.4506 | -101.1093
5.4543 | -101.1053
5.458 | -101.1203
5.4616 | -101.1514
5.4653 | -101.1981
5.4689 | -101.2472
5.4726 | -101.3055
5.4762 | -101.3524
5.4799 | -101.3863
5.4836 | -101.4018
5.4872 | -101.3963
5.4909 | -101.3726
5.4945 | -101.3339
5.4982 | -101.2893
5.5019 | -101.2431
5.5055 | -101.1924
5.5092 | -101.1449
5.5128 | -101.107
5.5165 | -101.0769
5.5201 | -101.061
5.5238 | -101.0552
5.5275 | -101.0522
5.5311 | -101.0627
5.5348 | -101.0738
5.5384 | -101.0877
5.5421 | -101.107
5.5457 | -101.1197
5.5494 | -101.1226
5.5531 | -101.1264
5.5567 | -101.1155
5.5604 | -101.1077
5.564 | -101.0979
5.5677 | -101.0963
5.5714 | -101.1073
5.575 | -101.1354
5.5787 | -101.1779
5.5823 | -101.2345
5.586 | -101.3059
5.5896 | -101.393
5.5933 | -101.4826
5.597 | -101.5636
5.6006 | -101.634
5.6043 | -101.6809
5.6079 | -101.7064
5.6116 | -101.7168
5.6153 | -101.7115
5.6189 | -101.6998
5.6226 | -101.6858
5.6262 | -101.6721
5.6299 | -101.6597
5.6335 | -101.6624
5.6372 | -101.6655
5.6409 | -101.6755
5.6445 | -101.676
5.6482 | -101.6708
5.6518 | -101.6537
5.6555 | -101.625
5.6592 | -101.5887
5.6628 | -101.5489
5.6665 | -101.5003
5.6701 | -101.4597
5.6738 | -101.434
5.6774 | -101.4157
5.6811 | -101.4109
5.6848 | -101.4173
5.6884 | -101.4349
5.6921 | -101.4535
5.6957 | -101.4682
5.6994 | -101.4717
5.7031 | -101.4641
5.7067 | -101.4351
5.7104 | -101.3967
5.714 | -101.3633
5.7177 | -101.3209
5.7213 | -101.2826
5.725 | -101.2476
5.7287 | -101.2234
5.7323 | -101.2054
5.736 | -101.2026
5.7396 | -101.2065
5.7433 | -101.2121
5.7469 | -101.229
5.7506 | -101.2501
5.7543 | -101.2707
5.7579 | -101.2965
5.7616 | -101.3263
5.7652 | -101.3582
5.7689 | -101.3884
5.7726 | -101.4189
5.7762 | -101.4473
5.7799 | -101.4724
5.7835 | -101.4974
5.7872 | -101.5204
5.7908 | -101.5457
5.7945 | -101.5679
5.7982 | -101.5843
5.8018 | -101.627
5.8055 | -101.6619
5.8091 | -101.7015
5.8128 | -101.7505
5.8165 | -101.8064
5.8201 | -101.8652
5.8238 | -101.913
5.8274 | -101.9676
5.8311 | -102.0262
5.8347 | -102.0725
5.8384 | -102.1163
5.8421 | -102.1504
5.8457 | -102.1771
5.8494 | -102.2001
5.853 | -102.216
5.8567 | -102.2224
5.8604 | -102.2215
5.864 | -102.2096
5.8677 | -102.1789
5.8713 | -102.1387
5.875 | -102.0897
5.8786 | -102.0316
5.8823 | -101.9714
5.886 | -101.9141
5.8896 | -101.8643
5.8933 | -101.8192
5.8969 | -101.7827
5.9006 | -101.7549
5.9042 | -101.7349
5.9079 | -101.7169
5.9116 | -101.7088
5.9152 | -101.7033
5.9189 | -101.6979
5.9225 | -101.6937
5.9262 | -101.6911
5.9299 | -101.6838
5.9335 | -101.6845
5.9372 | -101.6818
5.9408 | -101.6781
5.9445 | -101.6706
5.9481 | -101.6615
5.9518 | -101.6486
5.9555 | -101.638
5.9591 | -101.6323
5.9628 | -101.6251
5.9664 | -101.6202
5.9701 | -101.6224
5.9738 | -101.6297
5.9774 | -101.6494
5.9811 | -101.6795
5.9847 | -101.7216
5.9884 | -101.7749
5.992 | -101.8384
5.9957 | -101.9018
5.9994 | -101.9746
6.003 | -102.048
6.0067 | -102.1218
6.0103 | -102.1865
6.014 | -102.2497
6.0177 | -102.305
6.0213 | -102.3409
6.025 | -102.3536
6.0286 | -102.3596
6.0323 | -102.3464
6.0359 | -102.3179
6.0396 | -102.2781
6.0433 | -102.2321
6.0469 | -102.1742
6.0506 | -102.1043
6.0542 | -102.0305
6.0579 | -101.9582
6.0615 | -101.8755
6.0652 | -101.7974
6.0689 | -101.721
6.0725 | -101.6489
6.0762 | -101.5902
6.0798 | -101.5479
6.0835 | -101.5166
6.0872 | -101.4923
6.0908 | -101.4776
6.0945 | -101.4743
6.0981 | -101.4792
6.1018 | -101.4892
6.1054 | -101.4957
6.1091 | -101.5045
6.1128 | -101.5214
6.1164 | -101.531
6.1201 | -101.5419
6.1237 | -101.5645
6.1274 | -101.5899
6.1311 | -101.6183
6.1347 | -101.6526
6.1384 | -101.7011
6.142 | -101.7537
6.1457 | -101.8096
6.1493 | -101.8627
6.153 | -101.9138
6.1567 | -101.9486
6.1603 | -101.9797
6.164 | -102.0051
6.1676 | -102.0222
6.1713 | -102.031
6.175 | -102.0333
6.1786 | -102.0343
6.1823 | -102.0319
6.1859 | -102.0267
6.1896 | -102.0237
6.1932 | -102.0168
6.1969 | -102.011
6.2006 | -102.0062
6.2042 | -101.9995
6.2079 | -101.9986
6.2115 | -101.9933
6.2152 | -101.9847
6.2188 | -101.9773
6.2225 | -101.9744
6.2262 | -101.9652
6.2298 | -101.9592
6.2335 | -101.9656
6.2371 | -101.9756
6.2408 | -101.9925
6.2445 | -102.0231
6.2481 | -102.0577
6.2518 | -102.0931
6.2554 | -102.1322
6.2591 | -102.1677
6.2627 | -102.1998
6.2664 | -102.225
6.2701 | -102.2438
6.2737 | -102.2488
6.2774 | -102.2472
6.281 | -102.2347
6.2847 | -102.2179
6.2884 | -102.1814
6.292 | -102.1361
6.2957 | -102.096
6.2993 | -102.0506
6.303 | -102.0065
6.3066 | -101.9646
6.3103 | -101.9295
6.314 | -101.9008
6.3176 | -101.8823
6.3213 | -101.8746
6.3249 | -101.8815
6.3286 | -101.9076
6.3323 | -101.9807
6.3359 | -102.0336
6.3396 | -102.0937
6.3432 | -102.1629
6.3469 | -102.2314
6.3505 | -102.2947
6.3542 | -102.3551
6.3579 | -102.4054
6.3615 | -102.4474
6.3652 | -102.4731
6.3688 | -102.4856
6.3725 | -102.4849
6.3761 | -102.4735
6.3798 | -102.4516
6.3835 | -102.4373
6.3871 | -102.4134
6.3908 | -102.3727
6.3944 | -102.3337
6.3981 | -102.2933
6.4018 | -102.2559
6.4054 | -102.2243
6.4091 | -102.1961
6.4127 | -102.1786
6.4164 | -102.165
6.42 | -102.16
6.4237 | -102.1664
6.4274 | -102.1778
6.431 | -102.1936
6.4347 | -102.2215
6.4383 | -102.2489
6.442 | -102.2701
6.4457 | -102.2831
6.4493 | -102.2968
6.453 | -102.3017
6.4566 | -102.2953
6.4603 | -102.2789
6.4639 | -102.2563
6.4676 | -102.2298
6.4713 | -102.2051
6.4749 | -102.187
6.4786 | -102.1735
6.4822 | -102.1588
6.4859 | -102.1542
6.4896 | -102.1616
6.4932 | -102.1749
6.4969 | -102.1906
6.5005 | -102.2103
6.5042 | -102.2266
6.5078 | -102.2438
6.5115 | -102.2605
6.5152 | -102.2716
6.5188 | -102.2771
6.5225 | -102.2719
6.5261 | -102.2614
6.5298 | -102.242
6.5334 | -102.2101
6.5371 | -102.1758
6.5408 | -102.1415
6.5444 | -102.1049
6.5481 | -102.0702
6.5517 | -102.0461
6.5554 | -102.035
6.5591 | -102.0373
6.5627 | -102.0526
6.5664 | -102.0775
6.57 | -102.1175
6.5737 | -102.1622
6.5773 | -102.2132
6.581 | -102.2735
6.5847 | -102.3272
6.5883 | -102.3796
6.592 | -102.4296
6.5956 | -102.4713
6.5993 | -102.499
6.603 | -102.5135
6.6066 | -102.5147
6.6103 | -102.5061
6.6139 | -102.4927
6.6176 | -102.4757
6.6212 | -102.4639
6.6249 | -102.4543
6.6286 | -102.4513
6.6322 | -102.4555
6.6359 | -102.468
6.6395 | -102.4893
6.6432 | -102.5172
6.6469 | -102.5468
6.6505 | -102.5693
6.6542 | -102.5865
6.6578 | -102.5971
6.6615 | -102.5961
6.6651 | -102.5886
6.6688 | -102.5687
6.6725 | -102.5449
6.6761 | -102.516
6.6798 | -102.4863
6.6834 | -102.4666
6.6871 | -102.4509
6.6907 | -102.4453
6.6944 | -102.4523
6.6981 | -102.4714
6.7017 | -102.504
6.7054 | -102.5472
6.709 | -102.593
6.7127 | -102.6424
6.7164 | -102.6865
6.72 | -102.7192
6.7237 | -102.7515
6.7273 | -102.7733
6.731 | -102.7865
6.7346 | -102.7873
6.7383 | -102.7717
6.742 | -102.7497
6.7456 | -102.7275
6.7493 | -102.7165
6.7529 | -102.7046
6.7566 | -102.6941
6.7603 | -102.6971
6.7639 | -102.7082
6.7676 | -102.7231
6.7712 | -102.7405
6.7749 | -102.7584
6.7785 | -102.7708
6.7822 | -102.7816
6.7859 | -102.7831
6.7895 | -102.7783
6.7932 | -102.7674
6.7968 | -102.7461
6.8005 | -102.7197
6.8042 | -102.6882
6.8078 | -102.6468
6.8115 | -102.6025
6.8151 | -102.5588
6.8188 | -102.513
6.8224 | -102.4639
6.8261 | -102.417
6.8298 | -102.3788
6.8334 | -102.3538
6.8371 | -102.3504
6.8407 | -102.3585
6.8444 | -102.3755
6.848 | -102.4042
6.8517 | -102.4446
6.8554 | -102.4914
6.859 | -102.5256
6.8627 | -102.5527
6.8663 | -102.5733
6.87 | -102.5853
6.8737 | -102.5938
6.8773 | -102.5949
6.881 | -102.5908
6.8846 | -102.5783
6.8883 | -102.5654
6.8919 | -102.5512
6.8956 | -102.5383
6.8993 | -102.5285
6.9029 | -102.5173
6.9066 | -102.5057
6.9102 | -102.5006
6.9139 | -102.5003
6.9176 | -102.5026
6.9212 | -102.5113
6.9249 | -102.5216
6.9285 | -102.5384
6.9322 | -102.5588
6.9358 | -102.576
6.9395 | -102.5872
6.9432 | -102.6018
6.9468 | -102.6118
6.9505 | -102.6215
6.9541 | -102.63
6.9578 | -102.635
6.9615 | -102.6372
6.9651 | -102.6371
6.9688 | -102.6361
6.9724 | -102.6324
6.9761 | -102.6325
6.9797 | -102.632
6.9834 | -102.6341
6.9871 | -102.6352
6.9907 | -102.6375
6.9944 | -102.6444
6.998 | -102.6498
7.0017 | -102.6572
7.0053 | -102.6686
7.009 | -102.6831
7.0127 | -102.7024
7.0163 | -102.7284
7.02 | -102.7544
7.0236 | -102.7804
7.0273 | -102.807
7.031 | -102.8391
7.0346 | -102.8634
7.0383 | -102.881
7.0419 | -102.9025
7.0456 | -102.9163
7.0492 | -102.9208
7.0529 | -102.9193
7.0566 | -102.9154
7.0602 | -102.9087
7.0639 | -102.8983
7.0675 | -102.8894
7.0712 | -102.8812
7.0749 | -102.8674
7.0785 | -102.862
7.0822 | -102.8581
7.0858 | -102.8505
7.0895 | -102.8386
7.0931 | -102.8276
7.0968 | -102.8179
7.1005 | -102.7979
7.1041 | -102.7719
7.1078 | -102.7411
7.1114 | -102.7043
7.1151 | -102.6642
7.1188 | -102.6256
7.1224 | -102.5851
7.1261 | -102.5429
7.1297 | -102.5127
7.1334 | -102.4871
7.137 | -102.4685
7.1407 | -102.4589
7.1444 | -102.4541
7.148 | -102.4516
7.1517 | -102.4598
7.1553 | -102.4757
7.159 | -102.4969
7.1627 | -102.5178
7.1663 | -102.5477
7.17 | -102.5858
7.1736 | -102.627
7.1773 | -102.6759
7.1809 | -102.7268
7.1846 | -102.7707
7.1883 | -102.8149
7.1919 | -102.8641
7.1956 | -102.9128
7.1992 | -102.9553
7.2029 | -102.9814
7.2065 | -103.004
7.2102 | -103.0169
7.2139 | -103.0253
7.2175 | -103.0264
7.2212 | -103.0225
7.2248 | -103.0112
7.2285 | -103.0014
7.2322 | -102.9969
7.2358 | -102.9866
7.2395 | -102.9788
7.2431 | -102.9683
7.2468 | -102.9662
7.2504 | -102.9673
7.2541 | -102.9709
7.2578 | -102.9765
7.2614 | -102.9875
7.2651 | -103.0017
7.2687 | -103.0128
7.2724 | -103.0256
7.2761 | -103.0431
7.2797 | -103.0645
7.2834 | -103.0865
7.287 | -103.1123
7.2907 | -103.1385
7.2943 | -103.1635
7.298 | -103.1828
7.3017 | -103.205
7.3053 | -103.2309
7.309 | -103.2608
7.3126 | -103.286
7.3163 | -103.3109
7.32 | -103.326
7.3236 | -103.3373
7.3273 | -103.3422
7.3309 | -103.3452
7.3346 | -103.3439
7.3382 | -103.3334
7.3419 | -103.3197
7.3456 | -103.3016
7.3492 | -103.2885
7.3529 | -103.2767
7.3565 | -103.267
7.3602 | -103.2612
7.3638 | -103.2558
7.3675 | -103.2509
7.3712 | -103.2534
7.3748 | -103.2626
7.3785 | -103.266
7.3821 | -103.2627
7.3858 | -103.2551
7.3895 | -103.2448
7.3931 | -103.2335
7.3968 | -103.2176
7.4004 | -103.2006
7.4041 | -103.1842
7.4077 | -103.1681
7.4114 | -103.1447
7.4151 | -103.1178
7.4187 | -103.0899
7.4224 | -103.0586
7.426 | -103.0254
7.4297 | -102.9921
7.4334 | -102.96
7.437 | -102.9259
7.4407 | -102.8888
7.4443 | -102.8521
7.448 | -102.8257
7.4516 | -102.8062
7.4553 | -102.7878
7.459 | -102.7713
7.4626 | -102.7607
7.4663 | -102.7539
7.4699 | -102.7517
7.4736 | -102.7568
7.4773 | -102.7723
7.4809 | -102.7929
7.4846 | -102.8156
7.4882 | -102.8466
7.4919 | -102.8883
7.4955 | -102.9313
7.4992 | -102.9792
7.5029 | -103.0301
7.5065 | -103.078
7.5102 | -103.1296
7.5138 | -103.1777
7.5175 | -103.2226
7.5211 | -103.2644
7.5248 | -103.3001
7.5285 | -103.332
7.5321 | -103.3615
7.5358 | -103.3849
7.5394 | -103.4061
7.5431 | -103.4192
7.5468 | -103.428
7.5504 | -103.437
7.5541 | -103.4417
7.5577 | -103.4429
7.5614 | -103.4352
7.565 | -103.4231
7.5687 | -103.415
7.5724 | -103.4029
7.576 | -103.3889
7.5797 | -103.3766
7.5833 | -103.3678
7.587 | -103.3632
7.5907 | -103.3598
7.5943 | -103.3591
7.598 | -103.3654
7.6016 | -103.3691
7.6053 | -103.3679
7.6089 | -103.3701
7.6126 | -103.372
7.6163 | -103.3751
7.6199 | -103.3772
7.6236 | -103.3753
7.6272 | -103.3682
7.6309 | -103.3643
7.6346 | -103.3633
7.6382 | -103.3632
7.6419 | -103.3652
7.6455 | -103.3677
7.6492 | -103.3797
7.6528 | -103.3937
7.6565 | -103.4138
7.6602 | -103.4338
7.6638 | -103.455
7.6675 | -103.4772
7.6711 | -103.4966
7.6748 | -103.5211
7.6784 | -103.5502
7.6821 | -103.5701
7.6858 | -103.5891
7.6894 | -103.6037
7.6931 | -103.6126
7.6967 | -103.6141
7.7004 | -103.6083
7.7041 | -103.5966
7.7077 | -103.5787
7.7114 | -103.5588
7.715 | -103.537
7.7187 | -103.5086
7.7223 | -103.4761
7.726 | -103.4432
7.7297 | -103.4091
7.7333 | -103.379
7.737 | -103.3553
7.7406 | -103.3366
7.7443 | -103.3195
7.748 | -103.3095
7.7516 | -103.3113
7.7553 | -103.3143
7.7589 | -103.3197
7.7626 | -103.329
7.7662 | -103.3436
7.7699 | -103.3593
7.7736 | -103.3797
7.7772 | -103.4
7.7809 | -103.4221
7.7845 | -103.4469
7.7882 | -103.465
7.7919 | -103.4784
7.7955 | -103.5004
7.7992 | -103.5214
7.8028 | -103.5433
7.8065 | -103.5636
7.8101 | -103.5808
7.8138 | -103.5942
7.8175 | -103.6024
7.8211 | -103.6119
7.8248 | -103.6124
7.8284 | -103.6036
7.8321 | -103.5945
7.8357 | -103.577
7.8394 | -103.5519
7.8431 | -103.5248
7.8467 | -103.495
7.8504 | -103.4635
7.854 | -103.4359
7.8577 | -103.4085
7.8614 | -103.3779
7.865 | -103.3449
7.8687 | -103.3113
7.8723 | -103.2833
7.876 | -103.2526
7.8796 | -103.2167
7.8833 | -103.1897
7.887 | -103.1704
7.8906 | -103.1562
7.8943 | -103.1537
7.8979 | -103.1548
7.9016 | -103.1677
7.9053 | -103.1934
7.9089 | -103.2256
7.9126 | -103.2627
7.9162 | -103.3152
7.9199 | -103.3546
7.9235 | -103.3995
7.9272 | -103.4444
7.9309 | -103.4921
7.9345 | -103.5396
7.9382 | -103.5795
7.9418 | -103.6192
7.9455 | -103.6595
7.9492 | -103.6886
7.9528 | -103.709
7.9565 | -103.7252
7.9601 | -103.7329
7.9638 | -103.737
7.9674 | -103.7312
7.9711 | -103.7154
7.9748 | -103.6917
7.9784 | -103.6701
7.9821 | -103.6525
7.9857 | -103.6356
7.9894 | -103.6144
7.9931 | -103.6013
7.9967 | -103.5972
8.0004 | -103.5973
8.004 | -103.602
8.0077 | -103.6168
8.0113 | -103.6386
8.015 | -103.6697
8.0187 | -103.7048
8.0223 | -103.7404
8.026 | -103.7794
8.0296 | -103.8247
8.0333 | -103.8697
8.0369 | -103.908
8.0406 | -103.9433
8.0443 | -103.9709
8.0479 | -103.9935
8.0516 | -104.0104
8.0552 | -104.0235
8.0589 | -104.0314
8.0626 | -104.0287
8.0662 | -104.013
8.0699 | -103.997
8.0735 | -103.9807
8.0772 | -103.9547
8.0808 | -103.9221
8.0845 | -103.8878
8.0882 | -103.8577
8.0918 | -103.8296
8.0955 | -103.7999
8.0991 | -103.7668
8.1028 | -103.7307
8.1065 | -103.6984
8.1101 | -103.6695
8.1138 | -103.6421
8.1174 | -103.6128
8.1211 | -103.5848
8.1247 | -103.5589
8.1284 | -103.531
8.1321 | -103.5047
8.1357 | -103.4838
8.1394 | -103.4658
8.143 | -103.4533
8.1467 | -103.4442
8.1503 | -103.438
8.154 | -103.4356
8.1577 | -103.4428
8.1613 | -103.452
8.165 | -103.4607
8.1686 | -103.4753
8.1723 | -103.4916
8.176 | -103.5094
8.1796 | -103.5283
8.1833 | -103.5491
8.1869 | -103.5724
8.1906 | -103.5902
8.1942 | -103.6023
8.1979 | -103.6134
8.2016 | -103.6225
8.2052 | -103.6264
8.2089 | -103.6266
8.2125 | -103.6274
8.2162 | -103.6275
8.2199 | -103.6257
8.2235 | -103.6226
8.2272 | -103.6196
8.2308 | -103.6156
8.2345 | -103.6083
8.2381 | -103.6053
8.2418 | -103.6005
8.2455 | -103.596
8.2491 | -103.5943
8.2528 | -103.5866
8.2564 | -103.5769
8.2601 | -103.567
8.2638 | -103.5555
8.2674 | -103.5411
8.2711 | -103.5273
8.2747 | -103.5185
8.2784 | -103.5091
8.282 | -103.5022
8.2857 | -103.5008
8.2894 | -103.498
8.293 | -103.4962
8.2967 | -103.5018
8.3003 | -103.5155
8.304 | -103.5351
8.3077 | -103.561
8.3113 | -103.5878
8.315 | -103.617
8.3186 | -103.6487
8.3223 | -103.6865
8.3259 | -103.7286
8.3296 | -103.7708
8.3333 | -103.8114
8.3369 | -103.856
8.3406 | -103.8977
8.3442 | -103.9403
8.3479 | -103.9763
8.3515 | -104.0078
8.3552 | -104.0343
8.3589 | -104.057
8.3625 | -104.0725
8.3662 | -104.0865
8.3698 | -104.0972
8.3735 | -104.1042
8.3772 | -104.1069
8.3808 | -104.1009
8.3845 | -104.0971
8.3881 | -104.1009
8.3918 | -104.0993
8.3954 | -104.0944
8.3991 | -104.0905
8.4028 | -104.0881
8.4064 | -104.0821
8.4101 | -104.0755
8.4137 | -104.0677
8.4174 | -104.055
8.4211 | -104.0419
8.4247 | -104.0299
8.4284 | -104.0175
8.432 | -104.0161
8.4357 | -104.0184
8.4393 | -104.023
8.443 | -104.0395
8.4467 | -104.0467
8.4503 | -104.0508
8.454 | -104.0607
8.4576 | -104.0714
8.4613 | -104.0832
8.4649 | -104.0991
8.4686 | -104.1117
8.4723 | -104.1247
8.4759 | -104.1327
8.4796 | -104.1374
8.4832 | -104.1391
8.4869 | -104.1395
8.4906 | -104.1352
8.4942 | -104.1272
8.4979 | -104.1204
8.5015 | -104.1138
8.5052 | -104.1056
8.5088 | -104.0956
8.5125 | -104.0818
8.5162 | -104.0717
8.5198 | -104.067
8.5235 | -104.0635
8.5271 | -104.0589
8.5308 | -104.0575
8.5345 | -104.0573
8.5381 | -104.059
8.5418 | -104.0576
8.5454 | -104.0573
8.5491 | -104.0567
8.5527 | -104.0585
8.5564 | -104.0568
8.5601 | -104.0517
8.5637 | -104.0478
8.5674 | -104.046
8.571 | -104.045
8.5747 | -104.0459
8.5784 | -104.05
8.582 | -104.0491
8.5857 | -104.0514
8.5893 | -104.0531
8.593 | -104.0529
8.5966 | -104.0603
8.6003 | -104.0727
8.604 | -104.0904
8.6076 | -104.113
8.6113 | -104.1389
8.6149 | -104.1644
8.6186 | -104.189
8.6223 | -104.2175
8.6259 | -104.2426
8.6296 | -104.266
8.6332 | -104.2867
8.6369 | -104.3055
8.6405 | -104.3253
8.6442 | -104.3451
8.6479 | -104.3596
8.6515 | -104.3727
8.6552 | -104.3877
8.6588 | -104.4009
8.6625 | -104.4055
8.6661 | -104.4061
8.6698 | -104.4086
8.6735 | -104.4034
8.6771 | -104.3935
8.6808 | -104.3778
8.6844 | -104.3569
8.6881 | -104.3355
8.6918 | -104.3105
8.6954 | -104.2839
8.6991 | -104.2558
8.7027 | -104.224
8.7064 | -104.1912
8.71 | -104.1577
8.7137 | -104.119
8.7174 | -104.0774
8.721 | -104.0443
8.7247 | -104.0098
8.7283 | -103.9712
8.732 | -103.9327
8.7357 | -103.8922
8.7393 | -103.8529
8.743 | -103.82
8.7466 | -103.7911
8.7503 | -103.7653
8.7539 | -103.746
8.7576 | -103.7272
8.7613 | -103.7158
8.7649 | -103.7053
8.7686 | -103.697
8.7722 | -103.6947
8.7759 | -103.7009
8.7796 | -103.7069
8.7832 | -103.7167
8.7869 | -103.733
8.7905 | -103.7543
8.7942 | -103.7784
8.7978 | -103.8019
8.8015 | -103.8291
8.8052 | -103.8578
8.8088 | -103.8874
8.8125 | -103.9158
8.8161 | -103.9431
8.8198 | -103.9635
8.8234 | -103.9759
8.8271 | -103.9872
8.8308 | -103.9947
8.8344 | -104.0006
8.8381 | -104.0046
8.8417 | -104.0025
8.8454 | -103.9994
8.8491 | -103.9928
8.8527 | -103.9879
8.8564 | -103.987
8.86 | -103.9863
8.8637 | -103.9874
8.8673 | -103.9877
8.871 | -103.9922
8.8747 | -104.0032
8.8783 | -104.0118
8.882 | -104.0205
8.8856 | -104.0316
8.8893 | -104.0472
8.893 | -104.0659
8.8966 | -104.0872
8.9003 | -104.1058
8.9039 | -104.1223
8.9076 | -104.1475
8.9112 | -104.1722
8.9149 | -104.1877
8.9186 | -104.2032
8.9222 | -104.2143
8.9259 | -104.2166
8.9295 | -104.2131
8.9332 | -104.2087
8.9369 | -104.2004
8.9405 | -104.1868
8.9442 | -104.1673
8.9478 | -104.1441
8.9515 | -104.1201
8.9551 | -104.0945
8.9588 | -104.07
8.9625 | -104.0403
8.9661 | -104.0104
8.9698 | -103.9908
8.9734 | -103.9738
8.9771 | -103.9575
8.9807 | -103.9401
8.9844 | -103.9244
8.9881 | -103.9176
8.9917 | -103.915
8.9954 | -103.9171
8.999 | -103.926
9.0027 | -103.9357
9.0064 | -103.9443
9.01 | -103.9622
9.0137 | -103.9852
9.0173 | -104.0108
9.021 | -104.0336
9.0246 | -104.0561
9.0283 | -104.0799
9.032 | -104.1083
9.0356 | -104.1334
9.0393 | -104.1576
9.0429 | -104.1838
9.0466 | -104.2127
9.0503 | -104.2367
9.0539 | -104.256
9.0576 | -104.2758
9.0612 | -104.2927
9.0649 | -104.303
9.0685 | -104.3092
9.0722 | -104.3146
9.0759 | -104.3216
9.0795 | -104.3304
9.0832 | -104.3413
9.0868 | -104.3504
9.0905 | -104.3643
9.0942 | -104.3827
9.0978 | -104.4055
9.1015 | -104.4354
9.1051 | -104.4628
9.1088 | -104.4917
9.1124 | -104.5266
9.1161 | -104.5655
9.1198 | -104.6077
9.1234 | -104.6433
9.1271 | -104.6803
9.1307 | -104.714
9.1344 | -104.7418
9.138 | -104.7661
9.1417 | -104.7943
9.1454 | -104.8157
9.149 | -104.8305
9.1527 | -104.8444
9.1563 | -104.8543
9.16 | -104.8573
9.1637 | -104.8517
9.1673 | -104.8439
9.171 | -104.833
9.1746 | -104.8119
9.1783 | -104.787
9.1819 | -104.7603
9.1856 | -104.7329
9.1893 | -104.7019
9.1929 | -104.6743
9.1966 | -104.6414
9.2002 | -104.6109
9.2039 | -104.5862
9.2076 | -104.563
9.2112 | -104.536
9.2149 | -104.5119
9.2185 | -104.4915
9.2222 | -104.4705
9.2258 | -104.4496
9.2295 | -104.432
9.2332 | -104.4143
9.2368 | -104.3945
9.2405 | -104.3859
9.2441 | -104.381
9.2478 | -104.3756
9.2515 | -104.373
9.2551 | -104.3714
9.2588 | -104.3741
9.2624 | -104.3725
9.2661 | -104.3712
9.2697 | -104.3701
9.2734 | -104.3715
9.2771 | -104.3679
9.2807 | -104.3621
9.2844 | -104.3556
9.288 | -104.3464
9.2917 | -104.337
9.2953 | -104.3272
9.299 | -104.3217
9.3027 | -104.3082
9.3063 | -104.297
9.31 | -104.2881
9.3136 | -104.279
9.3173 | -104.2715
9.321 | -104.2639
9.3246 | -104.2553
9.3283 | -104.2429
9.3319 | -104.2389
9.3356 | -104.2374
9.3392 | -104.2329
9.3429 | -104.2262
9.3466 | -104.2269
9.3502 | -104.2289
9.3539 | -104.2281
9.3575 | -104.2309
9.3612 | -104.2337
9.3649 | -104.2381
9.3685 | -104.2432
9.3722 | -104.2483
9.3758 | -104.2496
9.3795 | -104.2523
9.3831 | -104.2527
9.3868 | -104.2509
9.3905 | -104.2521
9.3941 | -104.253
9.3978 | -104.2528
9.4014 | -104.2502
9.4051 | -104.2452
9.4088 | -104.2376
9.4124 | -104.2305
9.4161 | -104.2244
9.4197 | -104.22
9.4234 | -104.2142
9.427 | -104.2127
9.4307 | -104.2154
9.4344 | -104.215
9.438 | -104.218
9.4417 | -104.2243
9.4453 | -104.2286
9.449 | -104.2336
9.4527 | -104.2436
9.4563 | -104.2555
9.46 | -104.2704
9.4636 | -104.2822
9.4673 | -104.2916
9.4709 | -104.3017
9.4746 | -104.3086
9.4783 | -104.3217
9.4819 | -104.3336
9.4856 | -104.3439
9.4892 | -104.3543
9.4929 | -104.3666
9.4965 | -104.3817
9.5002 | -104.3937
9.5039 | -104.4031
9.5075 | -104.4114
9.5112 | -104.4225
9.5148 | -104.4326
9.5185 | -104.4389
9.5222 | -104.4481
9.5258 | -104.457
9.5295 | -104.4596
9.5331 | -104.4555
9.5368 | -104.4546
9.5404 | -104.4544
9.5441 | -104.4553
9.5478 | -104.4561
9.5514 | -104.4558
9.5551 | -104.4575
9.5587 | -104.4606
9.5624 | -104.4651
9.5661 | -104.4681
9.5697 | -104.4691
9.5734 | -104.4755
9.577 | -104.4865
9.5807 | -104.4942
9.5843 | -104.507
9.588 | -104.5184
9.5917 | -104.5233
9.5953 | -104.5312
9.599 | -104.5331
9.6026 | -104.533
9.6063 | -104.5295
9.6099 | -104.5249
9.6136 | -104.5232
9.6173 | -104.5178
9.6209 | -104.5068
9.6246 | -104.4962
9.6282 | -104.4879
9.6319 | -104.471
9.6356 | -104.4523
9.6392 | -104.4379
9.6429 | -104.4244
9.6465 | -104.4141
9.6502 | -104.4053
9.6538 | -104.3956
9.6575 | -104.3886
9.6612 | -104.3892
9.6648 | -104.3929
9.6685 | -104.3979
9.6721 | -104.4098
9.6758 | -104.4298
9.6795 | -104.4463
9.6831 | -104.4643
9.6868 | -104.4813
9.6904 | -104.4954
9.6941 | -104.5127
9.6977 | -104.5304
9.7014 | -104.5456
9.7051 | -104.5579
9.7087 | -104.5653
9.7124 | -104.5698
9.716 | -104.569
9.7197 | -104.5649
9.7234 | -104.5601
9.727 | -104.549
9.7307 | -104.535
9.7343 | -104.5196
9.738 | -104.4994
9.7416 | -104.4864
9.7453 | -104.4727
9.749 | -104.4587
9.7526 | -104.4451
9.7563 | -104.4357
9.7599 | -104.4274
9.7636 | -104.4233
9.7673 | -104.4242
9.7709 | -104.4289
9.7746 | -104.4399
9.7782 | -104.4493
9.7819 | -104.4584
9.7855 | -104.4708
9.7892 | -104.4855
9.7929 | -104.5046
9.7965 | -104.5268
9.8002 | -104.5465
9.8038 | -104.5641
9.8075 | -104.5798
9.8111 | -104.5948
9.8148 | -104.6086
9.8185 | -104.6196
9.8221 | -104.6312
9.8258 | -104.6385
9.8294 | -104.6432
9.8331 | -104.6506
9.8368 | -104.6488
9.8404 | -104.6402
9.8441 | -104.6278
9.8477 | -104.6185
9.8514 | -104.61
9.855 | -104.5995
9.8587 | -104.583
9.8624 | -104.5645
9.866 | -104.5498
9.8697 | -104.5391
9.8733 | -104.5276
9.877 | -104.5181
9.8807 | -104.5111
9.8843 | -104.4993
9.888 | -104.49
9.8916 | -104.4851
9.8953 | -104.48
9.8989 | -104.475
9.9026 | -104.4762
9.9063 | -104.4824
9.9099 | -104.4849
9.9136 | -104.4863
9.9172 | -104.4911
9.9209 | -104.4978
9.9246 | -104.5056
9.9282 | -104.5125
9.9319 | -104.5153
9.9355 | -104.5175
9.9392 | -104.521
9.9428 | -104.5274
9.9465 | -104.5323
9.9502 | -104.5351
9.9538 | -104.5352
9.9575 | -104.5291
9.9611 | -104.5183
9.9648 | -104.5087
9.9684 | -104.4992
9.9721 | -104.4902
9.9758 | -104.4778
9.9794 | -104.4639
9.9831 | -104.4549
9.9867 | -104.4439
9.9904 | -104.4313
9.9941 | -104.4226
9.9977 | -104.4201
10.0014 | -104.4216
10.005 | -104.4186
10.0087 | -104.4115
10.0123 | -104.4091
10.016 | -104.4117
10.0197 | -104.4157
10.0233 | -104.4198
10.027 | -104.4299
10.0306 | -104.4353
10.0343 | -104.4399
10.038 | -104.4476
10.0416 | -104.4559
10.0453 | -104.4637
10.0489 | -104.4717
10.0526 | -104.4818
10.0562 | -104.495
10.0599 | -104.5098
10.0636 | -104.5256
10.0672 | -104.5398
10.0709 | -104.5543
10.0745 | -104.5699
10.0782 | -104.5856
10.0819 | -104.5991
10.0855 | -104.6085
10.0892 | -104.6159
10.0928 | -104.6233
10.0965 | -104.6348
10.1001 | -104.6466
10.1038 | -104.6591
10.1075 | -104.6701
10.1111 | -104.6745
10.1148 | -104.6767
10.1184 | -104.6777
10.1221 | -104.6795
10.1257 | -104.6842
10.1294 | -104.6839
10.1331 | -104.6861
10.1367 | -104.686
10.1404 | -104.6825
10.144 | -104.6762
10.1477 | -104.6734
10.1514 | -104.6724
10.155 | -104.6733
10.1587 | -104.6711
10.1623 | -104.6698
10.166 | -104.6681
10.1696 | -104.6676
10.1733 | -104.6652
10.177 | -104.6635
10.1806 | -104.668
10.1843 | -104.6734
10.1879 | -104.6756
10.1916 | -104.6792
10.1953 | -104.6855
10.1989 | -104.6884
10.2026 | -104.6896
10.2062 | -104.695
10.2099 | -104.7005
10.2135 | -104.7057
10.2172 | -104.7133
10.2209 | -104.7147
10.2245 | -104.7128
10.2282 | -104.7143
10.2318 | -104.7151
10.2355 | -104.7167
10.2392 | -104.7137
10.2428 | -104.7057
10.2465 | -104.6925
10.2501 | -104.6792
10.2538 | -104.662
10.2574 | -104.6429
10.2611 | -104.6226
10.2648 | -104.601
10.2684 | -104.5763
10.2721 | -104.5474
10.2757 | -104.5198
10.2794 | -104.4922
10.283 | -104.4614
10.2867 | -104.435
10.2904 | -104.4114
10.294 | -104.3937
10.2977 | -104.3764
10.3013 | -104.3627
10.305 | -104.3494
10.3087 | -104.34
10.3123 | -104.3339
10.316 | -104.3345
10.3196 | -104.3385
10.3233 | -104.3449
10.3269 | -104.3475
10.3306 | -104.3515
10.3343 | -104.3605
10.3379 | -104.3752
10.3416 | -104.3952
10.3452 | -104.4135
10.3489 | -104.4328
10.3526 | -104.4549
10.3562 | -104.4746
10.3599 | -104.4916
10.3635 | -104.5015
10.3672 | -104.5109
10.3708 | -104.5237
10.3745 | -104.5316
10.3782 | -104.5339
10.3818 | -104.5327
10.3855 | -104.5267
10.3891 | -104.5233
10.3928 | -104.5154
10.3965 | -104.5054
10.4001 | -104.4947
10.4038 | -104.4857
10.4074 | -104.4783
10.4111 | -104.4712
10.4147 | -104.4631
10.4184 | -104.4595
10.4221 | -104.4594
10.4257 | -104.4626
10.4294 | -104.4686
10.433 | -104.4809
10.4367 | -104.4937
10.4403 | -104.5096
10.444 | -104.5263
10.4477 | -104.5507
10.4513 | -104.5833
10.455 | -104.6185
10.4586 | -104.6562
10.4623 | -104.6968
10.466 | -104.7356
10.4696 | -104.7772
10.4733 | -104.824
10.4769 | -104.8744
10.4806 | -104.9169
10.4842 | -104.9505
10.4879 | -104.9819
10.4916 | -105.0098
10.4952 | -105.0393
10.4989 | -105.067
10.5025 | -105.0879
10.5062 | -105.1069
10.5099 | -105.1268
10.5135 | -105.1473
10.5172 | -105.16
10.5208 | -105.1697
10.5245 | -105.176
10.5281 | -105.1821
10.5318 | -105.1865
10.5355 | -105.1902
10.5391 | -105.1882
10.5428 | -105.1899
10.5464 | -105.1901
10.5501 | -105.188
10.5538 | -105.1762
10.5574 | -105.171
10.5611 | -105.1623
10.5647 | -105.1523
10.5684 | -105.145
10.572 | -105.1333
10.5757 | -105.1217
10.5794 | -105.1068
10.583 | -105.0901
10.5867 | -105.0694
10.5903 | -105.0481
10.594 | -105.0287
10.5976 | -105.0091
10.6013 | -104.9848
10.605 | -104.9583
10.6086 | -104.9331
10.6123 | -104.9068
10.6159 | -104.8804
10.6196 | -104.8585
10.6233 | -104.8392
10.6269 | -104.8209
10.6306 | -104.8056
10.6342 | -104.795
10.6379 | -104.7872
10.6415 | -104.7755
10.6452 | -104.7671
10.6489 | -104.7603
10.6525 | -104.7563
10.6562 | -104.7558
10.6598 | -104.7522
10.6635 | -104.7526
10.6672 | -104.7555
10.6708 | -104.7603
10.6745 | -104.7675
10.6781 | -104.7806
10.6818 | -104.7909
10.6854 | -104.8037
10.6891 | -104.8219
10.6928 | -104.8401
10.6964 | -104.8569
10.7001 | -104.8773
10.7037 | -104.9005
10.7074 | -104.9264
10.7111 | -104.9473
10.7147 | -104.9689
10.7184 | -104.9938
10.722 | -105.0217
10.7257 | -105.0463
10.7293 | -105.071
10.733 | -105.0972
10.7367 | -105.1255
10.7403 | -105.1516
10.744 | -105.1764
10.7476 | -105.2004
10.7513 | -105.2224
10.7549 | -105.2439
10.7586 | -105.2643
10.7623 | -105.2803
10.7659 | -105.2948
10.7696 | -105.3071
10.7732 | -105.3129
10.7769 | -105.312
10.7806 | -105.3079
10.7842 | -105.3025
10.7879 | -105.2933
10.7915 | -105.2811
10.7952 | -105.2669
10.7988 | -105.2528
10.8025 | -105.236
10.8062 | -105.2156
10.8098 | -105.1961
10.8135 | -105.1785
10.8171 | -105.1632
10.8208 | -105.1426
10.8245 | -105.1261
10.8281 | -105.1114
10.8318 | -105.0942
10.8354 | -105.0831
10.8391 | -105.0774
10.8427 | -105.0719
10.8464 | -105.0662
10.8501 | -105.069
10.8537 | -105.0754
10.8574 | -105.0802
10.861 | -105.0865
10.8647 | -105.1021
10.8684 | -105.1203
10.872 | -105.137
10.8757 | -105.1525
10.8793 | -105.165
10.883 | -105.1791
10.8866 | -105.196
10.8903 | -105.2065
10.894 | -105.214
10.8976 | -105.2283
10.9013 | -105.2383
10.9049 | -105.2448
10.9086 | -105.251
10.9123 | -105.2542
10.9159 | -105.2581
10.9196 | -105.2615
10.9232 | -105.2621
10.9269 | -105.2614
10.9305 | -105.2618
10.9342 | -105.2611
10.9379 | -105.261
10.9415 | -105.2594
10.9452 | -105.2616
10.9488 | -105.2657
10.9525 | -105.2708
10.9561 | -105.2754
10.9598 | -105.2814
10.9635 | -105.2906
10.9671 | -105.2978
10.9708 | -105.3062
10.9744 | -105.3195
10.9781 | -105.3333
10.9818 | -105.3452
10.9854 | -105.3591
10.9891 | -105.3753
10.9927 | -105.3853
10.9964 | -105.3933
11 | -105.4035
11.0037 | -105.412
11.0074 | -105.4154
11.011 | -105.4145
11.0147 | -105.4114
11.0183 | -105.4036
11.022 | -105.3875
11.0257 | -105.3698
11.0293 | -105.3526
11.033 | -105.3369
11.0366 | -105.3221
11.0403 | -105.3051
11.0439 | -105.2824
11.0476 | -105.2632
11.0513 | -105.246
11.0549 | -105.224
11.0586 | -105.2066
11.0622 | -105.1913
11.0659 | -105.1774
11.0696 | -105.1654
11.0732 | -105.1515
11.0769 | -105.1393
11.0805 | -105.1291
11.0842 | -105.1162
11.0878 | -105.1016
11.0915 | -105.09
11.0952 | -105.0772
11.0988 | -105.0588
11.1025 | -105.0408
11.1061 | -105.0212
11.1098 | -105.0021
11.1134 | -104.9848
11.1171 | -104.9708
11.1208 | -104.9568
11.1244 | -104.9428
11.1281 | -104.9327
11.1317 | -104.9217
11.1354 | -104.9114
11.1391 | -104.9061
11.1427 | -104.9012
11.1464 | -104.8979
11.15 | -104.8995
11.1537 | -104.9014
11.1573 | -104.9059
11.161 | -104.9118
11.1647 | -104.92
11.1683 | -104.9295
11.172 | -104.9445
11.1756 | -104.9617
11.1793 | -104.9835
11.183 | -105.0072
11.1866 | -105.0301
11.1903 | -105.0511
11.1939 | -105.0718
11.1976 | -105.0872
11.2012 | -105.1077
11.2049 | -105.131
11.2086 | -105.1493
11.2122 | -105.1667
11.2159 | -105.1841
11.2195 | -105.197
11.2232 | -105.204
11.2269 | -105.2071
11.2305 | -105.2065
11.2342 | -105.203
11.2378 | -105.1956
11.2415 | -105.185
11.2451 | -105.1745
11.2488 | -105.1621
11.2525 | -105.1454
11.2561 | -105.1232
11.2598 | -105.0985
11.2634 | -105.0753
11.2671 | -105.052
11.2707 | -105.0292
11.2744 | -105.0073
11.2781 | -104.9863
11.2817 | -104.9657
11.2854 | -104.9473
11.289 | -104.9323
11.2927 | -104.9158
11.2964 | -104.9024
11.3 | -104.8919
11.3037 | -104.8787
11.3073 | -104.8714
11.311 | -104.8689
11.3146 | -104.8665
11.3183 | -104.8642
11.322 | -104.8647
11.3256 | -104.8661
11.3293 | -104.8665
11.3329 | -104.8698
11.3366 | -104.8732
11.3403 | -104.8768
11.3439 | -104.8789
11.3476 | -104.8811
11.3512 | -104.8824
11.3549 | -104.884
11.3585 | -104.8864
11.3622 | -104.8911
11.3659 | -104.8906
11.3695 | -104.8898
11.3732 | -104.8895
11.3768 | -104.8847
11.3805 | -104.8773
11.3842 | -104.8713
11.3878 | -104.8659
11.3915 | -104.8626
11.3951 | -104.863
11.3988 | -104.8627
11.4024 | -104.863
11.4061 | -104.8627
11.4098 | -104.8639
11.4134 | -104.8692
11.4171 | -104.8763
11.4207 | -104.8852
11.4244 | -104.8972
11.428 | -104.9095
11.4317 | -104.9193
11.4354 | -104.9328
11.439 | -104.9473
11.4427 | -104.9624
11.4463 | -104.9791
11.45 | -104.9951
11.4537 | -105.0119
11.4573 | -105.0312
11.461 | -105.0508
11.4646 | -105.0692
11.4683 | -105.0887
11.4719 | -105.1039
11.4756 | -105.1154
11.4793 | -105.1301
11.4829 | -105.1484
11.4866 | -105.1649
11.4902 | -105.1776
11.4939 | -105.1883
11.4976 | -105.1964
11.5012 | -105.2025
11.5049 | -105.2054
11.5085 | -105.2074
11.5122 | -105.2071
11.5158 | -105.208
11.5195 | -105.2066
11.5232 | -105.198
11.5268 | -105.1905
11.5305 | -105.1805
11.5341 | -105.1656
11.5378 | -105.1488
11.5415 | -105.1338
11.5451 | -105.1202
11.5488 | -105.1017
11.5524 | -105.0825
11.5561 | -105.0581
11.5597 | -105.0345
11.5634 | -105.0127
11.5671 | -104.9904
11.5707 | -104.9713
11.5744 | -104.9532
11.578 | -104.9388
11.5817 | -104.9267
11.5853 | -104.9136
11.589 | -104.9059
11.5927 | -104.9017
11.5963 | -104.9006
11.6 | -104.9011
11.6036 | -104.9084
11.6073 | -104.9301
11.611 | -104.941
11.6146 | -104.956
11.6183 | -104.9738
11.6219 | -104.9936
11.6256 | -105.0137
11.6292 | -105.0361
11.6329 | -105.0608
11.6366 | -105.086
11.6402 | -105.1093
11.6439 | -105.1342
11.6475 | -105.1546
11.6512 | -105.1724
11.6549 | -105.1826
11.6585 | -105.1954
11.6622 | -105.2068
11.6658 | -105.2135
11.6695 | -105.2159
11.6731 | -105.2191
11.6768 | -105.2174
11.6805 | -105.2116
11.6841 | -105.2018
11.6878 | -105.1913
11.6914 | -105.1801
11.6951 | -105.1627
11.6988 | -105.146
11.7024 | -105.1323
11.7061 | -105.1202
11.7097 | -105.1053
11.7134 | -105.0876
11.717 | -105.0709
11.7207 | -105.0567
11.7244 | -105.0469
11.728 | -105.0328
11.7317 | -105.02
11.7353 | -105.0121
11.739 | -105.0057
11.7426 | -105.0012
11.7463 | -104.9993
11.75 | -105.0012
11.7536 | -105.0055
11.7573 | -105.0083
11.7609 | -105.011
11.7646 | -105.0212
11.7683 | -105.0309
11.7719 | -105.0397
11.7756 | -105.0546
11.7792 | -105.0689
11.7829 | -105.0853
11.7865 | -105.1013
11.7902 | -105.1187
11.7939 | -105.1364
11.7975 | -105.1528
11.8012 | -105.169
11.8048 | -105.1848
11.8085 | -105.2021
11.8122 | -105.2193
11.8158 | -105.2351
11.8195 | -105.251
11.8231 | -105.2653
11.8268 | -105.285
11.8304 | -105.2996
11.8341 | -105.3126
11.8378 | -105.3249
11.8414 | -105.3363
11.8451 | -105.3465
11.8487 | -105.3544
11.8524 | -105.3618
11.8561 | -105.3686
11.8597 | -105.3725
11.8634 | -105.3774
11.867 | -105.386
11.8707 | -105.3905
11.8743 | -105.3924
11.878 | -105.3917
11.8817 | -105.3911
11.8853 | -105.391
11.889 | -105.3896
11.8926 | -105.3825
11.8963 | -105.3728
11.9 | -105.3672
11.9036 | -105.3616
11.9073 | -105.3536
11.9109 | -105.3436
11.9146 | -105.334
11.9182 | -105.3252
11.9219 | -105.3165
11.9256 | -105.308
11.9292 | -105.3018
11.9329 | -105.2969
11.9365 | -105.2869
11.9402 | -105.2772
11.9438 | -105.2664
11.9475 | -105.2541
11.9512 | -105.2462
11.9548 | -105.2364
11.9585 | -105.2242
11.9621 | -105.2124
11.9658 | -105.1992
11.9695 | -105.1895
11.9731 | -105.18
11.9768 | -105.1699
11.9804 | -105.16
11.9841 | -105.1519
11.9877 | -105.1431
11.9914 | -105.1317
11.9951 | -105.1271
11.9987 | -105.1213
12.0024 | -105.1152
12.006 | -105.1048
12.0097 | -105.1003
12.0134 | -105.0977
12.017 | -105.0941
12.0207 | -105.0924
12.0243 | -105.0909
12.028 | -105.0918
12.0316 | -105.0951
12.0353 | -105.0981
12.039 | -105.0971
12.0426 | -105.1009
12.0463 | -105.1056
12.0499 | -105.109
12.0536 | -105.1142
12.0572 | -105.1211
12.0609 | -105.1283
12.0646 | -105.1373
12.0682 | -105.1436
12.0719 | -105.1469
12.0755 | -105.1482
12.0792 | -105.1467
12.0829 | -105.1506
12.0865 | -105.1552
12.0902 | -105.1578
12.0938 | -105.1622
12.0975 | -105.1687
12.1011 | -105.1711
12.1048 | -105.1715
12.1085 | -105.1714
12.1121 | -105.1734
12.1158 | -105.1738
12.1194 | -105.1701
12.1231 | -105.1654
12.1268 | -105.159
12.1304 | -105.1539
12.1341 | -105.1457
12.1377 | -105.1245
12.1414 | -105.1133
12.145 | -105.0994
12.1487 | -105.0829
12.1524 | -105.0643
12.156 | -105.0468
12.1597 | -105.0293
12.1633 | -105.011
12.167 | -104.9953
12.1707 | -104.9768
12.1743 | -104.9577
12.178 | -104.9448
12.1816 | -104.9358
12.1853 | -104.923
12.1889 | -104.9114
12.1926 | -104.9029
12.1963 | -104.8996
12.1999 | -104.8985
12.2036 | -104.8954
12.2072 | -104.8938
12.2109 | -104.895
12.2145 | -104.8993
12.2182 | -104.9028
12.2219 | -104.9072
12.2255 | -104.9174
12.2292 | -104.9256
12.2328 | -104.9346
12.2365 | -104.9459
12.2402 | -104.9567
12.2438 | -104.9693
12.2475 | -104.9837
12.2511 | -104.9984
12.2548 | -105.0105
12.2584 | -105.0198
12.2621 | -105.0301
12.2658 | -105.042
12.2694 | -105.0553
12.2731 | -105.067
12.2767 | -105.0797
12.2804 | -105.0952
12.2841 | -105.111
12.2877 | -105.1229
12.2914 | -105.1365
12.295 | -105.1497
12.2987 | -105.1622
12.3023 | -105.1761
12.306 | -105.1863
12.3097 | -105.1983
12.3133 | -105.2153
12.317 | -105.2271
12.3206 | -105.2372
12.3243 | -105.2471
12.328 | -105.2572
12.3316 | -105.2684
12.3353 | -105.2781
12.3389 | -105.284
12.3426 | -105.2939
12.3462 | -105.3062
12.3499 | -105.3206
12.3536 | -105.3327
12.3572 | -105.3389
12.3609 | -105.3441
12.3645 | -105.3507
12.3682 | -105.359
12.3719 | -105.3657
12.3755 | -105.3669
12.3792 | -105.3651
12.3828 | -105.3661
12.3865 | -105.3677
12.3901 | -105.3649
12.3938 | -105.3618
12.3975 | -105.356
12.4011 | -105.3443
12.4048 | -105.3277
12.4084 | -105.3156
12.4121 | -105.3015
12.4157 | -105.2867
12.4194 | -105.272
12.4231 | -105.2585
12.4267 | -105.2444
12.4304 | -105.231
12.434 | -105.2182
12.4377 | -105.2036
12.4414 | -105.1894
12.445 | -105.1782
12.4487 | -105.1693
12.4523 | -105.1634
12.456 | -105.1576
12.4596 | -105.1539
12.4633 | -105.1544
12.467 | -105.1564
12.4706 | -105.1585
12.4743 | -105.1615
12.4779 | -105.1667
12.4816 | -105.1697
12.4853 | -105.1792
12.4889 | -105.1877
12.4926 | -105.1911
12.4962 | -105.2007
12.4999 | -105.2141
12.5035 | -105.2245
12.5072 | -105.2364
12.5109 | -105.2477
12.5145 | -105.2563
12.5182 | -105.2669
12.5218 | -105.2718
12.5255 | -105.2759
12.5292 | -105.2781
12.5328 | -105.2808
12.5365 | -105.2872
12.5401 | -105.2901
12.5438 | -105.2888
12.5474 | -105.2868
12.5511 | -105.2852
12.5548 | -105.2864
12.5584 | -105.2871
12.5621 | -105.2847
12.5657 | -105.2779
12.5694 | -105.2724
12.573 | -105.2686
12.5767 | -105.2684
12.5804 | -105.2702
12.584 | -105.2733
12.5877 | -105.2723
12.5913 | -105.2715
12.595 | -105.2703
12.5987 | -105.273
12.6023 | -105.2774
12.606 | -105.2813
12.6096 | -105.286
12.6133 | -105.2912
12.6169 | -105.2986
12.6206 | -105.3063
12.6243 | -105.3109
12.6279 | -105.3148
12.6316 | -105.3188
12.6352 | -105.3268
12.6389 | -105.3356
12.6426 | -105.3384
12.6462 | -105.3433
12.6499 | -105.3529
12.6535 | -105.3642
12.6572 | -105.3693
12.6608 | -105.3736
12.6645 | -105.3782
12.6682 | -105.3832
12.6718 | -105.389
12.6755 | -105.3958
12.6791 | -105.3962
12.6828 | -105.3962
12.6865 | -105.4
12.6901 | -105.4036
12.6938 | -105.4085
12.6974 | -105.4149
12.7011 | -105.4211
12.7047 | -105.4289
12.7084 | -105.4363
12.7121 | -105.4454
12.7157 | -105.4556
12.7194 | -105.4632
12.723 | -105.4761
12.7267 | -105.4892
12.7303 | -105.5008
12.734 | -105.5125
12.7377 | -105.5242
12.7413 | -105.5377
12.745 | -105.551
12.7486 | -105.5704
12.7523 | -105.591
12.756 | -105.6102
12.7596 | -105.6251
12.7633 | -105.6429
12.7669 | -105.6591
12.7706 | -105.6773
12.7742 | -105.7004
12.7779 | -105.7245
12.7816 | -105.7433
12.7852 | -105.7599
12.7889 | -105.7767
12.7925 | -105.7914
12.7962 | -105.8087
12.7999 | -105.8203
12.8035 | -105.8309
12.8072 | -105.8418
12.8108 | -105.8493
12.8145 | -105.8538
12.8181 | -105.8529
12.8218 | -105.8479
12.8255 | -105.8468
12.8291 | -105.8439
12.8328 | -105.8381
12.8364 | -105.8271
12.8401 | -105.8185
12.8438 | -105.8124
12.8474 | -105.8007
12.8511 | -105.7871
12.8547 | -105.7732
12.8584 | -105.7595
12.862 | -105.747
12.8657 | -105.7367
12.8694 | -105.726
12.873 | -105.7162
12.8767 | -105.7037
12.8803 | -105.6913
12.884 | -105.6797
12.8876 | -105.6673
12.8913 | -105.6573
12.895 | -105.647
12.8986 | -105.6361
12.9023 | -105.6274
12.9059 | -105.6184
12.9096 | -105.609
12.9133 | -105.6019
12.9169 | -105.5991
12.9206 | -105.595
12.9242 | -105.5926
12.9279 | -105.5929
12.9315 | -105.5959
12.9352 | -105.5921
12.9389 | -105.5892
12.9425 | -105.5904
12.9462 | -105.5955
12.9498 | -105.6009
12.9535 | -105.6088
12.9572 | -105.6163
12.9608 | -105.6226
12.9645 | -105.6298
12.9681 | -105.6376
12.9718 | -105.6469
12.9754 | -105.6595
12.9791 | -105.6709
12.9828 | -105.678
12.9864 | -105.6859
12.9901 | -105.6927
12.9937 | -105.7008
12.9974 | -105.7063
13.0011 | -105.7091
13.0047 | -105.7151
13.0084 | -105.7211
13.012 | -105.7267
13.0157 | -105.7328
13.0193 | -105.738
13.023 | -105.7446
13.0267 | -105.7453
13.0303 | -105.7451
13.034 | -105.7492
13.0376 | -105.7562
13.0413 | -105.7621
13.0449 | -105.767
13.0486 | -105.7691
13.0523 | -105.7712
13.0559 | -105.7725
13.0596 | -105.7713
13.0632 | -105.7707
13.0669 | -105.7727
13.0706 | -105.7737
13.0742 | -105.7779
13.0779 | -105.7794
13.0815 | -105.7791
13.0852 | -105.7781
13.0888 | -105.7792
13.0925 | -105.7812
13.0962 | -105.7795
13.0998 | -105.7757
13.1035 | -105.7679
13.1071 | -105.7619
13.1108 | -105.7495
13.1145 | -105.736
13.1181 | -105.7231
13.1218 | -105.7135
13.1254 | -105.7004
13.1291 | -105.6828
13.1327 | -105.6651
13.1364 | -105.645
13.1401 | -105.6227
13.1437 | -105.6015
13.1474 | -105.58
13.151 | -105.5571
13.1547 | -105.5376
13.1584 | -105.5173
13.162 | -105.4949
13.1657 | -105.4753
13.1693 | -105.4612
13.173 | -105.4479
13.1766 | -105.4344
13.1803 | -105.4206
13.184 | -105.4136
13.1876 | -105.4124
13.1913 | -105.4226
13.1949 | -105.4244
13.1986 | -105.4279
13.2022 | -105.4325
13.2059 | -105.4374
13.2096 | -105.4464
13.2132 | -105.4602
13.2169 | -105.4768
13.2205 | -105.4919
13.2242 | -105.5068
13.2279 | -105.5221
13.2315 | -105.5374
13.2352 | -105.5565
13.2388 | -105.5754
13.2425 | -105.5941
13.2461 | -105.6112
13.2498 | -105.6297
13.2535 | -105.6494
13.2571 | -105.6668
13.2608 | -105.6824
13.2644 | -105.6938
13.2681 | -105.7067
13.2718 | -105.718
13.2754 | -105.7278
13.2791 | -105.7336
13.2827 | -105.7343
13.2864 | -105.7351
13.29 | -105.7348
13.2937 | -105.7342
13.2974 | -105.7298
13.301 | -105.727
13.3047 | -105.7159
13.3083 | -105.7011
13.312 | -105.6825
13.3157 | -105.661
13.3193 | -105.6392
13.323 | -105.6147
13.3266 | -105.5907
13.3303 | -105.5676
13.3339 | -105.5422
13.3376 | -105.5131
13.3413 | -105.4819
13.3449 | -105.4481
13.3486 | -105.4168
13.3522 | -105.3907
13.3559 | -105.3628
13.3596 | -105.3334
13.3632 | -105.3034
13.3669 | -105.2748
13.3705 | -105.2489
13.3742 | -105.2224
13.3778 | -105.1956
13.3815 | -105.1692
13.3852 | -105.1414
13.3888 | -105.1196
13.3925 | -105.0974
13.3961 | -105.0779
13.3998 | -105.0578
13.4034 | -105.0392
13.4071 | -105.0247
13.4108 | -105.0142
13.4144 | -105.0058
13.4181 | -104.9961
13.4217 | -104.9888
13.4254 | -104.9852
13.4291 | -104.9832
13.4327 | -104.9839
13.4364 | -104.9861
13.44 | -104.991
13.4437 | -104.9965
13.4473 | -105.0018
13.451 | -105.0087
13.4547 | -105.0214
13.4583 | -105.0377
13.462 | -105.0529
13.4656 | -105.0703
13.4693 | -105.0891
13.473 | -105.1084
13.4766 | -105.1302
13.4803 | -105.1524
13.4839 | -105.176
13.4876 | -105.197
13.4912 | -105.2186
13.4949 | -105.2391
13.4986 | -105.2595
13.5022 | -105.2801
13.5059 | -105.3029
13.5095 | -105.3248
13.5132 | -105.3444
13.5168 | -105.3657
13.5205 | -105.3873
13.5242 | -105.407
13.5278 | -105.4269
13.5315 | -105.4459
13.5351 | -105.4641
13.5388 | -105.4853
13.5425 | -105.5037
13.5461 | -105.5169
13.5498 | -105.5343
13.5534 | -105.5533
13.5571 | -105.5709
13.5607 | -105.5843
13.5644 | -105.5962
13.5681 | -105.608
13.5717 | -105.6203
13.5754 | -105.6322
13.579 | -105.6424
13.5827 | -105.6448
13.5864 | -105.6517
13.59 | -105.6575
13.5937 | -105.6627
13.5973 | -105.6608
13.601 | -105.6592
13.6046 | -105.66
13.6083 | -105.6583
13.612 | -105.6586
13.6156 | -105.6555
13.6193 | -105.6529
13.6229 | -105.651
13.6266 | -105.6514
13.6303 | -105.6501
13.6339 | -105.6477
13.6376 | -105.6447
13.6412 | -105.6429
13.6449 | -105.6404
13.6485 | -105.6391
13.6522 | -105.6364
13.6559 | -105.6342
13.6595 | -105.6305
13.6632 | -105.6288
13.6668 | -105.6283
13.6705 | -105.6283
13.6742 | -105.6297
13.6778 | -105.6292
13.6815 | -105.6319
13.6851 | -105.6368
13.6888 | -105.6418
13.6924 | -105.6475
13.6961 | -105.6504
13.6998 | -105.652
13.7034 | -105.6526
13.7071 | -105.6538
13.7107 | -105.6536
13.7144 | -105.6523
13.718 | -105.646
13.7217 | -105.6427
13.7254 | -105.6397
13.729 | -105.6388
13.7327 | -105.6382
13.7363 | -105.6297
13.74 | -105.6207
13.7437 | -105.6119
13.7473 | -105.5988
13.751 | -105.5824
13.7546 | -105.5708
13.7583 | -105.5614
13.7619 | -105.55
13.7656 | -105.5365
13.7693 | -105.524
13.7729 | -105.513
13.7766 | -105.5024
13.7802 | -105.4932
13.7839 | -105.4795
13.7876 | -105.4674
13.7912 | -105.4601
13.7949 | -105.458
13.7985 | -105.4526
13.8022 | -105.4475
13.8058 | -105.4451
13.8095 | -105.4432
13.8132 | -105.444
13.8168 | -105.444
13.8205 | -105.4444
13.8241 | -105.4479
13.8278 | -105.4535
13.8315 | -105.462
13.8351 | -105.4716
13.8388 | -105.4794
13.8424 | -105.4886
13.8461 | -105.4947
13.8497 | -105.5017
13.8534 | -105.5107
13.8571 | -105.5201
13.8607 | -105.5282
13.8644 | -105.5363
13.868 | -105.5461
13.8717 | -105.5551
13.8753 | -105.5636
13.879 | -105.573
13.8827 | -105.5816
13.8863 | -105.5936
13.89 | -105.6067
13.8936 | -105.6196
13.8973 | -105.6327
13.901 | -105.6433
13.9046 | -105.6508
13.9083 | -105.6574
13.9119 | -105.6644
13.9156 | -105.6752
13.9192 | -105.6822
13.9229 | -105.6891
13.9266 | -105.6953
13.9302 | -105.7007
13.9339 | -105.7062
13.9375 | -105.7103
13.9412 | -105.7112
13.9449 | -105.7073
13.9485 | -105.7069
13.9522 | -105.7094
13.9558 | -105.7055
13.9595 | -105.6897
13.9631 | -105.6709
13.9668 | -105.6512
13.9705 | -105.6306
13.9741 | -105.6089
13.9778 | -105.5872
13.9814 | -105.5609
13.9851 | -105.5364
13.9888 | -105.5105
13.9924 | -105.4803
13.9961 | -105.4472
13.9997 | -105.4123
14.0034 | -105.374
14.007 | -105.3339
14.0107 | -105.2924
14.0144 | -105.2505
14.018 | -105.2089
14.0217 | -105.1679
14.0253 | -105.1252
14.029 | -105.0823
14.0326 | -105.0398
14.0363 | -104.9976
14.04 | -104.9614
14.0436 | -104.9288
14.0473 | -104.8981
14.0509 | -104.8728
14.0546 | -104.8525
14.0583 | -104.8367
14.0619 | -104.8236
14.0656 | -104.815
14.0692 | -104.8083
14.0729 | -104.808
14.0765 | -104.8146
14.0802 | -104.824
14.0839 | -104.8364
14.0875 | -104.8517
14.0912 | -104.8725
14.0948 | -104.8953
14.0985 | -104.9182
14.1022 | -104.9469
14.1058 | -104.9787
14.1095 | -105.01
14.1131 | -105.0416
14.1168 | -105.0705
14.1204 | -105.1022
14.1241 | -105.1363
14.1278 | -105.1682
14.1314 | -105.2013
14.1351 | -105.2296
14.1387 | -105.2589
14.1424 | -105.2895
14.1461 | -105.3218
14.1497 | -105.3498
14.1534 | -105.3755
14.157 | -105.4036
14.1607 | -105.4391
14.1643 | -105.4736
14.168 | -105.5027
14.1717 | -105.5278
14.1753 | -105.5503
14.179 | -105.57
14.1826 | -105.5905
14.1863 | -105.6085
14.1899 | -105.6231
14.1936 | -105.6378
14.1973 | -105.6541
14.2009 | -105.6705
14.2046 | -105.6879
14.2082 | -105.7066
14.2119 | -105.7251
14.2156 | -105.7414
14.2192 | -105.7571
14.2229 | -105.7687
14.2265 | -105.7827
14.2302 | -105.7979
14.2338 | -105.8117
14.2375 | -105.8256
14.2412 | -105.8355
14.2448 | -105.8433
14.2485 | -105.8429
14.2521 | -105.8492
14.2558 | -105.854
14.2595 | -105.8563
14.2631 | -105.8569
14.2668 | -105.8598
14.2704 | -105.8602
14.2741 | -105.8595
14.2777 | -105.8558
14.2814 | -105.853
14.2851 | -105.8485
14.2887 | -105.8452
14.2924 | -105.8377
14.296 | -105.8275
14.2997 | -105.8158
14.3034 | -105.8041
14.307 | -105.7924
14.3107 | -105.7801
14.3143 | -105.7664
14.318 | -105.7477
14.3216 | -105.7308
14.3253 | -105.7129
14.329 | -105.6933
14.3326 | -105.6735
14.3363 | -105.6502
14.3399 | -105.6287
14.3436 | -105.6056
14.3473 | -105.5807
14.3509 | -105.558
14.3546 | -105.5328
14.3582 | -105.5056
14.3619 | -105.482
14.3655 | -105.4581
14.3692 | -105.4328
14.3729 | -105.4127
14.3765 | -105.3942
14.3802 | -105.3765
14.3838 | -105.3593
14.3875 | -105.3431
14.3911 | -105.3298
14.3948 | -105.3197
14.3985 | -105.311
14.4021 | -105.3032
14.4058 | -105.2982
14.4094 | -105.296
14.4131 | -105.295
14.4168 | -105.294
14.4204 | -105.2968
14.4241 | -105.3005
14.4277 | -105.3073
14.4314 | -105.3137
14.435 | -105.3218
14.4387 | -105.3314
14.4424 | -105.3385
14.446 | -105.3489
14.4497 | -105.3588
14.4533 | -105.3687
14.457 | -105.3778
14.4607 | -105.386
14.4643 | -105.3924
14.468 | -105.394
14.4716 | -105.3992
14.4753 | -105.4009
14.4789 | -105.4053
14.4826 | -105.4088
14.4863 | -105.4071
14.4899 | -105.4009
14.4936 | -105.3966
14.4972 | -105.3911
14.5009 | -105.3879
14.5046 | -105.3842
14.5082 | -105.3798
14.5119 | -105.3769
14.5155 | -105.3706
14.5192 | -105.362
14.5228 | -105.3546
14.5265 | -105.3497
14.5302 | -105.3419
14.5338 | -105.333
14.5375 | -105.3263
14.5411 | -105.321
14.5448 | -105.3178
14.5484 | -105.3164
14.5521 | -105.3172
14.5558 | -105.3185
14.5594 | -105.3187
14.5631 | -105.3223
14.5667 | -105.3288
14.5704 | -105.3341
14.5741 | -105.3433
14.5777 | -105.3558
14.5814 | -105.3695
14.585 | -105.3843
14.5887 | -105.3995
14.5923 | -105.4182
14.596 | -105.4413
14.5997 | -105.4646
14.6033 | -105.4859
14.607 | -105.5097
14.6106 | -105.5352
14.6143 | -105.5599
14.618 | -105.5827
14.6216 | -105.6062
14.6253 | -105.6318
14.6289 | -105.6563
14.6326 | -105.6804
14.6362 | -105.7006
14.6399 | -105.7184
14.6436 | -105.7363
14.6472 | -105.7537
14.6509 | -105.773
14.6545 | -105.7866
14.6582 | -105.7998
14.6618 | -105.8144
14.6655 | -105.8292
14.6692 | -105.84
14.6728 | -105.8521
14.6765 | -105.866
14.6801 | -105.8761
14.6838 | -105.8873
14.6875 | -105.8973
14.6911 | -105.9068
14.6948 | -105.9191
14.6984 | -105.9308
14.7021 | -105.9414
14.7057 | -105.952
14.7094 | -105.9623
14.7131 | -105.9718
14.7167 | -105.9817
14.7204 | -105.9892
14.724 | -105.9972
14.7277 | -106.0084
14.7314 | -106.018
14.735 | -106.0245
14.7387 | -106.0327
14.7423 | -106.0428
14.746 | -106.0519
14.7496 | -106.0624
14.7533 | -106.0697
14.757 | -106.0715
14.7606 | -106.0732
14.7643 | -106.0771
14.7679 | -106.0812
14.7716 | -106.085
14.7753 | -106.0781
14.7789 | -106.0779
14.7826 | -106.0769
14.7862 | -106.0733
14.7899 | -106.0698
14.7935 | -106.0626
14.7972 | -106.0537
14.8009 | -106.048
14.8045 | -106.0403
14.8082 | -106.0248
14.8118 | -106.0081
14.8155 | -105.9947
14.8192 | -105.9782
14.8228 | -105.9584
14.8265 | -105.9409
14.8301 | -105.9251
14.8338 | -105.9098
14.8374 | -105.8942
14.8411 | -105.8797
14.8448 | -105.8647
14.8484 | -105.856
14.8521 | -105.8478
14.8557 | -105.8366
14.8594 | -105.8269
14.863 | -105.8179
14.8667 | -105.81
14.8704 | -105.8021
14.874 | -105.7979
14.8777 | -105.7953
14.8813 | -105.7918
14.885 | -105.7896
14.8887 | -105.7911
14.8923 | -105.7912
14.896 | -105.7901
14.8996 | -105.7863
14.9033 | -105.7869
14.9069 | -105.7904
14.9106 | -105.7935
14.9143 | -105.7964
14.9179 | -105.7945
14.9216 | -105.794
14.9252 | -105.7933
14.9289 | -105.7926
14.9326 | -105.792
14.9362 | -105.7913
14.9399 | -105.7873
14.9435 | -105.7865
14.9472 | -105.7887
14.9508 | -105.7849
14.9545 | -105.7774
14.9582 | -105.7693
14.9618 | -105.7591
14.9655 | -105.7515
14.9691 | -105.7448
14.9728 | -105.7321
14.9765 | -105.7185
14.9801 | -105.7016
14.9838 | -105.6839
14.9874 | -105.6647
14.9911 | -105.6429
14.9947 | -105.623
14.9984 | -105.6014
15.0021 | -105.5789
15.0057 | -105.5558
15.0094 | -105.5336
15.013 | -105.5108
15.0167 | -105.4859
15.0203 | -105.4599
15.024 | -105.4393
15.0277 | -105.4173
15.0313 | -105.3928
15.035 | -105.3669
15.0386 | -105.3444
15.0423 | -105.3224
15.046 | -105.3017
15.0496 | -105.2815
15.0533 | -105.2614
15.0569 | -105.2417
15.0606 | -105.2254
15.0642 | -105.2106
15.0679 | -105.2
15.0716 | -105.1933
15.0752 | -105.1889
15.0789 | -105.1878
15.0825 | -105.1883
15.0862 | -105.1909
15.0899 | -105.1968
15.0935 | -105.2068
15.0972 | -105.2211
15.1008 | -105.2379
15.1045 | -105.2565
15.1081 | -105.2786
15.1118 | -105.3049
15.1155 | -105.3314
15.1191 | -105.3595
15.1228 | -105.3878
15.1264 | -105.419
15.1301 | -105.4501
15.1338 | -105.4848
15.1374 | -105.5162
15.1411 | -105.5473
15.1447 | -105.5771
15.1484 | -105.6076
15.152 | -105.6361
15.1557 | -105.6612
15.1594 | -105.6831
15.163 | -105.7029
15.1667 | -105.7226
15.1703 | -105.7395
15.174 | -105.7522
15.1776 | -105.7626
15.1813 | -105.7692
15.185 | -105.7741
15.1886 | -105.7789
15.1923 | -105.7776
15.1959 | -105.7756
15.1996 | -105.7708
15.2033 | -105.7648
15.2069 | -105.758
15.2106 | -105.7476
15.2142 | -105.7347
15.2179 | -105.7243
15.2215 | -105.7143
15.2252 | -105.7038
15.2289 | -105.6969
15.2325 | -105.6914
15.2362 | -105.6875
15.2398 | -105.6804
15.2435 | -105.6755
15.2472 | -105.6731
15.2508 | -105.6759
15.2545 | -105.6769
15.2581 | -105.6787
15.2618 | -105.6834
15.2654 | -105.6912
15.2691 | -105.6994
15.2728 | -105.7083
15.2764 | -105.7164
15.2801 | -105.7245
15.2837 | -105.7333
15.2874 | -105.7431
15.2911 | -105.7515
15.2947 | -105.7611
15.2984 | -105.7703
15.302 | -105.7715
15.3057 | -105.7803
15.3093 | -105.785
15.313 | -105.7881
15.3167 | -105.7881
15.3203 | -105.7859
15.324 | -105.7838
15.3276 | -105.7844
15.3313 | -105.7826
15.3349 | -105.7831
15.3386 | -105.7791
15.3423 | -105.7736
15.3459 | -105.768
15.3496 | -105.7641
15.3532 | -105.7594
15.3569 | -105.7538
15.3606 | -105.7493
15.3642 | -105.7464
15.3679 | -105.7416
15.3715 | -105.737
15.3752 | -105.7336
15.3788 | -105.7302
15.3825 | -105.725
15.3862 | -105.7223
15.3898 | -105.7221
15.3935 | -105.7228
15.3971 | -105.7278
15.4008 | -105.7328
15.4045 | -105.7354
15.4081 | -105.7386
15.4118 | -105.7431
15.4154 | -105.7487
15.4191 | -105.7547
15.4227 | -105.7614
15.4264 | -105.7684
15.4301 | -105.7743
15.4337 | -105.7805
15.4374 | -105.7838
15.441 | -105.7842
15.4447 | -105.7792
15.4484 | -105.7751
15.452 | -105.7723
15.4557 | -105.7683
15.4593 | -105.7648
15.463 | -105.7611
15.4666 | -105.7559
15.4703 | -105.7448
15.474 | -105.732
15.4776 | -105.7194
15.4813 | -105.7056
15.4849 | -105.6873
15.4886 | -105.668
15.4922 | -105.6508
15.4959 | -105.633
15.4996 | -105.6113
15.5032 | -105.5847
15.5069 | -105.5616
15.5105 | -105.5367
15.5142 | -105.5096
15.5179 | -105.4835
15.5215 | -105.4544
15.5252 | -105.4255
15.5288 | -105.3956
15.5325 | -105.366
15.5361 | -105.3374
15.5398 | -105.314
15.5435 | -105.2895
15.5471 | -105.2643
15.5508 | -105.2368
15.5544 | -105.213
15.5581 | -105.1916
15.5618 | -105.1726
15.5654 | -105.1566
15.5691 | -105.1427
15.5727 | -105.1278
15.5764 | -105.1181
15.58 | -105.1111
15.5837 | -105.1079
15.5874 | -105.1067
15.591 | -105.1079
15.5947 | -105.1132
15.5983 | -105.1209
15.602 | -105.1293
15.6057 | -105.1402
15.6093 | -105.1548
15.613 | -105.1715
15.6166 | -105.1895
15.6203 | -105.2054
15.6239 | -105.2247
15.6276 | -105.2464
15.6313 | -105.2693
15.6349 | -105.2929
15.6386 | -105.3177
15.6422 | -105.34
15.6459 | -105.3603
15.6495 | -105.381
15.6532 | -105.4019
15.6569 | -105.4223
15.6605 | -105.4435
15.6642 | -105.4623
15.6678 | -105.4838
15.6715 | -105.5008
15.6752 | -105.5144
15.6788 | -105.5272
15.6825 | -105.5399
15.6861 | -105.5496
15.6898 | -105.5592
15.6934 | -105.568
15.6971 | -105.5745
15.7008 | -105.5766
15.7044 | -105.5759
15.7081 | -105.5779
15.7117 | -105.5788
15.7154 | -105.5766
15.7191 | -105.5725
15.7227 | -105.5693
15.7264 | -105.5684
15.73 | -105.563
15.7337 | -105.5597
15.7373 | -105.5562
15.741 | -105.5505
15.7447 | -105.5462
15.7483 | -105.5415
15.752 | -105.5393
15.7556 | -105.5362
15.7593 | -105.5324
15.763 | -105.5261
15.7666 | -105.5203
15.7703 | -105.5163
15.7739 | -105.5131
15.7776 | -105.509
15.7812 | -105.5022
15.7849 | -105.4983
15.7886 | -105.4954
15.7922 | -105.4925
15.7959 | -105.4917
15.7995 | -105.4882
15.8032 | -105.488
15.8069 | -105.488
15.8105 | -105.4898
15.8142 | -105.4913
15.8178 | -105.4914
15.8215 | -105.4944
15.8251 | -105.4993
15.8288 | -105.503
15.8325 | -105.5072
15.8361 | -105.5102
15.8398 | -105.511
15.8434 | -105.5138
15.8471 | -105.5164
15.8507 | -105.5177
15.8544 | -105.5211
15.8581 | -105.526
15.8617 | -105.5281
15.8654 | -105.5274
15.869 | -105.5308
15.8727 | -105.5356
15.8764 | -105.537
15.88 | -105.5383
15.8837 | -105.54
15.8873 | -105.5399
15.891 | -105.5405
15.8946 | -105.5426
15.8983 | -105.5441
15.902 | -105.5461
15.9056 | -105.5497
15.9093 | -105.5527
15.9129 | -105.5556
15.9166 | -105.5603
15.9203 | -105.5688
15.9239 | -105.5787
15.9276 | -105.5901
15.9312 | -105.6019
15.9349 | -105.6127
15.9385 | -105.6239
15.9422 | -105.636
15.9459 | -105.6512
15.9495 | -105.6685
15.9532 | -105.6861
15.9568 | -105.7022
15.9605 | -105.7191
15.9642 | -105.7386
15.9678 | -105.7599
15.9715 | -105.7804
15.9751 | -105.8002
15.9788 | -105.8181
15.9824 | -105.8359
15.9861 | -105.8547
15.9898 | -105.8734
15.9934 | -105.8922
15.9971 | -105.9089
16.0007 | -105.9238
16.0044 | -105.938
16.008 | -105.9506
16.0117 | -105.9596
16.0154 | -105.9654
16.019 | -105.9687
16.0227 | -105.9711
16.0263 | -105.9712
16.03 | -105.966
16.0337 | -105.9576
16.0373 | -105.9467
16.041 | -105.9354
16.0446 | -105.9236
16.0483 | -105.9078
16.0519 | -105.8898
16.0556 | -105.8709
16.0593 | -105.8492
16.0629 | -105.8264
16.0666 | -105.8025
16.0702 | -105.7779
16.0739 | -105.7529
16.0776 | -105.7303
16.0812 | -105.7062
16.0849 | -105.6787
16.0885 | -105.6484
16.0922 | -105.6245
16.0958 | -105.6034
16.0995 | -105.5798
16.1032 | -105.5552
16.1068 | -105.5343
16.1105 | -105.5161
16.1141 | -105.5002
16.1178 | -105.4832
16.1214 | -105.4646
16.1251 | -105.4487
16.1288 | -105.4398
16.1324 | -105.4292
16.1361 | -105.4182
16.1397 | -105.411
16.1434 | -105.4088
16.1471 | -105.4098
16.1507 | -105.4102
16.1544 | -105.413
16.158 | -105.417
16.1617 | -105.4217
16.1653 | -105.4288
16.169 | -105.4392
16.1727 | -105.4482
16.1763 | -105.4555
16.18 | -105.4621
16.1836 | -105.4719
16.1873 | -105.4828
16.191 | -105.4925
16.1946 | -105.5018
16.1983 | -105.5124
16.2019 | -105.5251
16.2056 | -105.5365
16.2092 | -105.5489
16.2129 | -105.5607
16.2166 | -105.5692
16.2202 | -105.5781
16.2239 | -105.5868
16.2275 | -105.5948
16.2312 | -105.6046
16.2349 | -105.6123
16.2385 | -105.62
16.2422 | -105.6273
16.2458 | -105.637
16.2495 | -105.6471
16.2531 | -105.6559
16.2568 | -105.6659
16.2605 | -105.6768
16.2641 | -105.6868
16.2678 | -105.6977
16.2714 | -105.7098
16.2751 | -105.7201
16.2788 | -105.7281
16.2824 | -105.7368
16.2861 | -105.7464
16.2897 | -105.7562
16.2934 | -105.7645
16.297 | -105.7732
16.3007 | -105.7846
16.3044 | -105.7936
16.308 | -105.8033
16.3117 | -105.8149
16.3153 | -105.8282
16.319 | -105.8409
16.3226 | -105.8522
16.3263 | -105.8626
16.33 | -105.8741
16.3336 | -105.8874
16.3373 | -105.8997
16.3409 | -105.9107
16.3446 | -105.9201
16.3483 | -105.9309
16.3519 | -105.9474
16.3556 | -105.9596
16.3592 | -105.9659
16.3629 | -105.9694
16.3665 | -105.9766
16.3702 | -105.9841
16.3739 | -105.9879
16.3775 | -105.9905
16.3812 | -105.9914
16.3848 | -105.993
16.3885 | -105.9938
16.3922 | -105.9943
16.3958 | -105.9952
16.3995 | -105.9942
16.4031 | -105.9915
16.4068 | -105.9899
16.4104 | -105.99
16.4141 | -105.9877
16.4178 | -105.9818
16.4214 | -105.976
16.4251 | -105.9701
16.4287 | -105.9618
16.4324 | -105.9528
16.4361 | -105.9424
16.4397 | -105.9317
16.4434 | -105.919
16.447 | -105.9093
16.4507 | -105.9017
16.4543 | -105.8924
16.458 | -105.8815
16.4617 | -105.8688
16.4653 | -105.8566
16.469 | -105.8448
16.4726 | -105.836
16.4763 | -105.8251
16.4799 | -105.8109
16.4836 | -105.7951
16.4873 | -105.7806
16.4909 | -105.7683
16.4946 | -105.7575
16.4982 | -105.7471
16.5019 | -105.7384
16.5056 | -105.7292
16.5092 | -105.722
16.5129 | -105.7199
16.5165 | -105.7141
16.5202 | -105.7067
16.5238 | -105.7007
16.5275 | -105.6941
16.5312 | -105.6912
16.5348 | -105.6883
16.5385 | -105.685
16.5421 | -105.6815
16.5458 | -105.6786
16.5495 | -105.6747
16.5531 | -105.671
16.5568 | -105.6658
16.5604 | -105.6632
16.5641 | -105.6607
16.5677 | -105.6568
16.5714 | -105.654
16.5751 | -105.652
16.5787 | -105.6493
16.5824 | -105.6459
16.586 | -105.6423
16.5897 | -105.6393
16.5934 | -105.6378
16.597 | -105.6362
16.6007 | -105.6333
16.6043 | -105.6293
16.608 | -105.6254
16.6116 | -105.6232
16.6153 | -105.6195
16.619 | -105.6171
16.6226 | -105.6151
16.6263 | -105.6142
16.6299 | -105.6111
16.6336 | -105.609
16.6372 | -105.6072
16.6409 | -105.6079
16.6446 | -105.6082
16.6482 | -105.6068
16.6519 | -105.6085
16.6555 | -105.6114
16.6592 | -105.6146
16.6629 | -105.6206
16.6665 | -105.6287
16.6702 | -105.6359
16.6738 | -105.6426
16.6775 | -105.6504
16.6811 | -105.6597
16.6848 | -105.6673
16.6885 | -105.674
16.6921 | -105.6809
16.6958 | -105.6892
16.6994 | -105.6997
16.7031 | -105.7117
16.7068 | -105.7226
16.7104 | -105.7305
16.7141 | -105.7407
16.7177 | -105.7533
16.7214 | -105.7674
16.725 | -105.7803
16.7287 | -105.7895
16.7324 | -105.799
16.736 | -105.809
16.7397 | -105.8153
16.7433 | -105.8222
16.747 | -105.8284
16.7507 | -105.8305
16.7543 | -105.8349
16.758 | -105.8412
16.7616 | -105.845
16.7653 | -105.8489
16.7689 | -105.8469
16.7726 | -105.8421
16.7763 | -105.8367
16.7799 | -105.8312
16.7836 | -105.8273
16.7872 | -105.8228
16.7909 | -105.8176
16.7945 | -105.812
16.7982 | -105.8044
16.8019 | -105.7952
16.8055 | -105.7866
16.8092 | -105.7786
16.8128 | -105.7679
16.8165 | -105.7592
16.8202 | -105.752
16.8238 | -105.7427
16.8275 | -105.7336
16.8311 | -105.7259
16.8348 | -105.7216
16.8384 | -105.7219
16.8421 | -105.7216
16.8458 | -105.722
16.8494 | -105.7208
16.8531 | -105.7209
16.8567 | -105.7221
16.8604 | -105.7261
16.8641 | -105.7313
16.8677 | -105.7383
16.8714 | -105.7445
16.875 | -105.7526
16.8787 | -105.7591
16.8823 | -105.7684
16.886 | -105.7833
16.8897 | -105.7938
16.8933 | -105.8024
16.897 | -105.8126
16.9006 | -105.8247
16.9043 | -105.8356
16.908 | -105.8453
16.9116 | -105.8547
16.9153 | -105.8661
16.9189 | -105.8769
16.9226 | -105.8855
16.9262 | -105.8958
16.9299 | -105.9047
16.9336 | -105.9103
16.9372 | -105.9168
16.9409 | -105.9235
16.9445 | -105.9263
16.9482 | -105.9308
16.9519 | -105.9378
16.9555 | -105.9417
16.9592 | -105.945
16.9628 | -105.9492
16.9665 | -105.9529
16.9701 | -105.9555
16.9738 | -105.9555
16.9775 | -105.9551
16.9811 | -105.9569
16.9848 | -105.959
16.9884 | -105.9621
16.9921 | -105.9669
16.9957 | -105.9696
16.9994 | -105.9714
17.0031 | -105.9724
17.0067 | -105.9743
17.0104 | -105.9798
17.014 | -105.9869
17.0177 | -105.9939
17.0214 | -105.9984
17.025 | -106.001
17.0287 | -106.0077
17.0323 | -106.0138
17.036 | -106.0181
17.0396 | -106.023
17.0433 | -106.0297
17.047 | -106.0368
17.0506 | -106.0425
17.0543 | -106.0495
17.0579 | -106.0561
17.0616 | -106.0613
17.0653 | -106.0679
17.0689 | -106.074
17.0726 | -106.077
17.0762 | -106.0796
17.0799 | -106.0815
17.0835 | -106.0873
17.0872 | -106.0912
17.0909 | -106.0955
17.0945 | -106.1001
17.0982 | -106.1022
17.1018 | -106.1042
17.1055 | -106.1045
17.1091 | -106.1076
17.1128 | -106.1124
17.1165 | -106.1172
17.1201 | -106.1193
17.1238 | -106.1197
17.1274 | -106.1218
17.1311 | -106.1214
17.1348 | -106.1205
17.1384 | -106.1212
17.1421 | -106.1208
17.1457 | -106.1204
17.1494 | -106.12
17.153 | -106.1167
17.1567 | -106.1155
17.1604 | -106.1138
17.164 | -106.1116
17.1677 | -106.1066
17.1713 | -106.1008
17.175 | -106.0979
17.1787 | -106.0984
17.1823 | -106.0937
17.186 | -106.0861
17.1896 | -106.0786
17.1933 | -106.0718
17.1969 | -106.0636
17.2006 | -106.0553
17.2043 | -106.0441
17.2079 | -106.0332
17.2116 | -106.02
17.2152 | -106.0082
17.2189 | -106
17.2226 | -105.9907
17.2262 | -105.9798
17.2299 | -105.9678
17.2335 | -105.9531
17.2372 | -105.9394
17.2408 | -105.9276
17.2445 | -105.9142
17.2482 | -105.9019
17.2518 | -105.8885
17.2555 | -105.8691
17.2591 | -105.8525
17.2628 | -105.8384
17.2664 | -105.8235
17.2701 | -105.8073
17.2738 | -105.7919
17.2774 | -105.7754
17.2811 | -105.7599
17.2847 | -105.7455
17.2884 | -105.7308
17.2921 | -105.7128
17.2957 | -105.6945
17.2994 | -105.6785
17.303 | -105.6641
17.3067 | -105.6504
17.3103 | -105.6342
17.314 | -105.6168
17.3177 | -105.6038
17.3213 | -105.5923
17.325 | -105.5793
17.3286 | -105.5647
17.3323 | -105.5545
17.336 | -105.5454
17.3396 | -105.5355
17.3433 | -105.5283
17.3469 | -105.521
17.3506 | -105.5133
17.3542 | -105.5064
17.3579 | -105.4991
17.3616 | -105.4908
17.3652 | -105.4866
17.3689 | -105.4828
17.3725 | -105.4788
17.3762 | -105.4802
17.3799 | -105.484
17.3835 | -105.4868
17.3872 | -105.4909
17.3908 | -105.4935
17.3945 | -105.4983
17.3981 | -105.5074
17.4018 | -105.5165
17.4055 | -105.5256
17.4091 | -105.5367
17.4128 | -105.5547
17.4164 | -105.5677
17.4201 | -105.5802
17.4238 | -105.5932
17.4274 | -105.6057
17.4311 | -105.6203
17.4347 | -105.6361
17.4384 | -105.653
17.442 | -105.668
17.4457 | -105.6847
17.4494 | -105.7
17.453 | -105.7158
17.4567 | -105.7313
17.4603 | -105.7477
17.464 | -105.7656
17.4676 | -105.7787
17.4713 | -105.7914
17.475 | -105.8045
17.4786 | -105.8161
17.4823 | -105.8292
17.4859 | -105.8402
17.4896 | -105.8473
17.4933 | -105.8534
17.4969 | -105.8581
17.5006 | -105.8638
17.5042 | -105.8719
17.5079 | -105.8762
17.5115 | -105.8786
17.5152 | -105.8809
17.5189 | -105.8826
17.5225 | -105.8786
17.5262 | -105.875
17.5298 | -105.8714
17.5335 | -105.8653
17.5372 | -105.8599
17.5408 | -105.8558
17.5445 | -105.8498
17.5481 | -105.8415
17.5518 | -105.8311
17.5554 | -105.8217
17.5591 | -105.8109
17.5628 | -105.8006
17.5664 | -105.7884
17.5701 | -105.7757
17.5737 | -105.7651
17.5774 | -105.7532
17.5811 | -105.74
17.5847 | -105.7286
17.5884 | -105.7198
17.592 | -105.7107
17.5957 | -105.7024
17.5993 | -105.6966
17.603 | -105.6899
17.6067 | -105.6838
17.6103 | -105.6811
17.614 | -105.6789
17.6176 | -105.6775
17.6213 | -105.6752
17.6249 | -105.6713
17.6286 | -105.6702
17.6323 | -105.6726
17.6359 | -105.6766
17.6396 | -105.6842
17.6432 | -105.692
17.6469 | -105.6998
17.6506 | -105.7061
17.6542 | -105.7157
17.6579 | -105.7251
17.6615 | -105.7346
17.6652 | -105.7461
17.6688 | -105.7615
17.6725 | -105.7763
17.6762 | -105.7905
17.6798 | -105.8044
17.6835 | -105.8174
17.6871 | -105.832
17.6908 | -105.846
17.6945 | -105.8602
17.6981 | -105.8735
17.7018 | -105.8888
17.7054 | -105.9035
17.7091 | -105.9176
17.7127 | -105.9307
17.7164 | -105.9397
17.7201 | -105.9488
17.7237 | -105.9579
17.7274 | -105.9662
17.731 | -105.9761
17.7347 | -105.9834
17.7384 | -105.9903
17.742 | -105.9991
17.7457 | -106.0069
17.7493 | -106.0135
17.753 | -106.0164
17.7566 | -106.0194
17.7603 | -106.0227
17.764 | -106.0266
17.7676 | -106.0312
17.7713 | -106.0341
17.7749 | -106.0376
17.7786 | -106.0429
17.7822 | -106.0445
17.7859 | -106.0458
17.7896 | -106.0461
17.7932 | -106.048
17.7969 | -106.0517
17.8005 | -106.0558
17.8042 | -106.059
17.8079 | -106.0583
17.8115 | -106.0598
17.8152 | -106.0631
17.8188 | -106.064
17.8225 | -106.0646
17.8261 | -106.0637
17.8298 | -106.063
17.8335 | -106.0637
17.8371 | -106.0595
17.8408 | -106.0545
17.8444 | -106.0509
17.8481 | -106.0451
17.8518 | -106.0397
17.8554 | -106.0337
17.8591 | -106.0259
17.8627 | -106.0149
17.8664 | -106.0028
17.87 | -105.9888
17.8737 | -105.9753
17.8774 | -105.9607
17.881 | -105.9442
17.8847 | -105.9256
17.8883 | -105.9054
17.892 | -105.8838
17.8957 | -105.8643
17.8993 | -105.8434
17.903 | -105.8252
17.9066 | -105.8073
17.9103 | -105.7901
17.9139 | -105.7734
17.9176 | -105.7569
17.9213 | -105.7388
17.9249 | -105.7224
17.9286 | -105.7055
17.9322 | -105.6891
17.9359 | -105.6767
17.9395 | -105.6633
17.9432 | -105.6587
17.9469 | -105.6495
17.9505 | -105.6373
17.9542 | -105.6255
17.9578 | -105.6152
17.9615 | -105.6042
17.9652 | -105.5947
17.9688 | -105.5867
17.9725 | -105.5796
17.9761 | -105.5737
17.9798 | -105.57
17.9834 | -105.5666
17.9871 | -105.5636
17.9908 | -105.5633
17.9944 | -105.5622
17.9981 | -105.5606
18.0017 | -105.5596
18.0054 | -105.5602
18.0091 | -105.5614
18.0127 | -105.5635
18.0164 | -105.5638
18.02 | -105.5628
18.0237 | -105.5624
18.0273 | -105.5638
18.031 | -105.5667
18.0347 | -105.5687
18.0383 | -105.5716
18.042 | -105.5734
18.0456 | -105.574
18.0493 | -105.5746
18.053 | -105.5758
18.0566 | -105.5756
18.0603 | -105.5761
18.0639 | -105.5801
18.0676 | -105.5825
18.0712 | -105.5834
18.0749 | -105.5871
18.0786 | -105.5895
18.0822 | -105.5943
18.0859 | -105.6007
18.0895 | -105.6037
18.0932 | -105.6105
18.0968 | -105.6177
18.1005 | -105.623
18.1042 | -105.6294
18.1078 | -105.636
18.1115 | -105.643
18.1151 | -105.6502
18.1188 | -105.6575
18.1225 | -105.6644
18.1261 | -105.672
18.1298 | -105.6784
18.1334 | -105.6851
18.1371 | -105.6906
18.1407 | -105.695
18.1444 | -105.6992
18.1481 | -105.7045
18.1517 | -105.707
18.1554 | -105.7097
18.159 | -105.7118
18.1627 | -105.7136
18.1664 | -105.7168
18.17 | -105.719
18.1737 | -105.7201
18.1773 | -105.7203
18.181 | -105.7216
18.1846 | -105.7244
18.1883 | -105.7264
18.192 | -105.7265
18.1956 | -105.7275
18.1993 | -105.7307
18.2029 | -105.7325
18.2066 | -105.7332
18.2103 | -105.7396
18.2139 | -105.742
18.2176 | -105.7434
18.2212 | -105.7461
18.2249 | -105.7466
18.2285 | -105.7501
18.2322 | -105.7519
18.2359 | -105.7524
18.2395 | -105.7532
18.2432 | -105.7541
18.2468 | -105.754
18.2505 | -105.7529
18.2541 | -105.7524
18.2578 | -105.7531
18.2615 | -105.7552
18.2651 | -105.7568
18.2688 | -105.7599
18.2724 | -105.7619
18.2761 | -105.7655
18.2798 | -105.7712
18.2834 | -105.7754
18.2871 | -105.7789
18.2907 | -105.7826
18.2944 | -105.7867
18.298 | -105.7914
18.3017 | -105.798
18.3054 | -105.8062
18.309 | -105.8108
18.3127 | -105.8141
18.3163 | -105.8215
18.32 | -105.83
18.3237 | -105.8378
18.3273 | -105.8456
18.331 | -105.8546
18.3346 | -105.8621
18.3383 | -105.8721
18.3419 | -105.8805
18.3456 | -105.8884
18.3493 | -105.8975
18.3529 | -105.9063
18.3566 | -105.9174
18.3602 | -105.9297
18.3639 | -105.9408
18.3676 | -105.9498
18.3712 | -105.9591
18.3749 | -105.9712
18.3785 | -105.9838
18.3822 | -105.9961
18.3858 | -106.0052
18.3895 | -106.0157
18.3932 | -106.026
18.3968 | -106.0352
18.4005 | -106.0453
18.4041 | -106.0545
18.4078 | -106.0615
18.4114 | -106.0693
18.4151 | -106.0765
18.4188 | -106.0823
18.4224 | -106.0855
18.4261 | -106.0884
18.4297 | -106.093
18.4334 | -106.0984
18.4371 | -106.1044
18.4407 | -106.1088
18.4444 | -106.1116
18.448 | -106.1146
18.4517 | -106.1166
18.4553 | -106.1164
18.459 | -106.1154
18.4627 | -106.1141
18.4663 | -106.1133
18.47 | -106.1129
18.4736 | -106.1125
18.4773 | -106.1135
18.481 | -106.1141
18.4846 | -106.1148
18.4883 | -106.1144
18.4919 | -106.1111
18.4956 | -106.1088
18.4992 | -106.1071
18.5029 | -106.1034
18.5066 | -106.0993
18.5102 | -106.0971
18.5139 | -106.0972
18.5175 | -106.0991
18.5212 | -106.0981
18.5249 | -106.0966
18.5285 | -106.0941
18.5322 | -106.092
18.5358 | -106.0903
18.5395 | -106.0904
18.5431 | -106.0884
18.5468 | -106.0866
18.5505 | -106.0853
18.5541 | -106.0813
18.5578 | -106.0764
18.5614 | -106.0709
18.5651 | -106.0649
18.5688 | -106.0598
18.5724 | -106.057
18.5761 | -106.0548
18.5797 | -106.0476
18.5834 | -106.0394
18.587 | -106.0324
18.5907 | -106.0252
18.5944 | -106.0148
18.598 | -106.0086
18.6017 | -106.0029
18.6053 | -105.9967
18.609 | -105.9887
18.6126 | -105.9816
18.6163 | -105.9742
18.62 | -105.9625
18.6236 | -105.9525
18.6273 | -105.9459
18.6309 | -105.9397
18.6346 | -105.9326
18.6383 | -105.9242
18.6419 | -105.9171
18.6456 | -105.9086
18.6492 | -105.903
18.6529 | -105.8949
18.6565 | -105.8851
18.6602 | -105.8778
18.6639 | -105.8705
18.6675 | -105.863
18.6712 | -105.8556
18.6748 | -105.8468
18.6785 | -105.8416
18.6822 | -105.8372
18.6858 | -105.8338
18.6895 | -105.8302
18.6931 | -105.8243
18.6968 | -105.8205
18.7004 | -105.8185
18.7041 | -105.8169
18.7078 | -105.817
18.7114 | -105.8194
18.7151 | -105.8214
18.7187 | -105.8263
18.7224 | -105.8306
18.7261 | -105.836
18.7297 | -105.8432
18.7334 | -105.8494
18.737 | -105.8546
18.7407 | -105.8613
18.7443 | -105.8703
18.748 | -105.8798
18.7517 | -105.8882
18.7553 | -105.8952
18.759 | -105.9048
18.7626 | -105.9152
18.7663 | -105.9247
18.7699 | -105.9345
18.7736 | -105.9442
18.7773 | -105.9537
18.7809 | -105.9603
18.7846 | -105.9667
18.7882 | -105.9762
18.7919 | -105.9849
18.7956 | -105.9932
18.7992 | -106.0018
18.8029 | -106.0096
18.8065 | -106.0165
18.8102 | -106.0248
18.8138 | -106.034
18.8175 | -106.0425
18.8212 | -106.0518
18.8248 | -106.0628
18.8285 | -106.0744
18.8321 | -106.0833
18.8358 | -106.0887
18.8395 | -106.0935
18.8431 | -106.0974
18.8468 | -106.1043
18.8504 | -106.112
18.8541 | -106.1169
18.8577 | -106.1213
18.8614 | -106.1272
18.8651 | -106.127
18.8687 | -106.1261
18.8724 | -106.129
18.876 | -106.134
18.8797 | -106.1358
18.8834 | -106.1367
18.887 | -106.1391
18.8907 | -106.1412
18.8943 | -106.1428
18.898 | -106.1444
18.9016 | -106.1452
18.9053 | -106.1463
18.909 | -106.1469
18.9126 | -106.1481
18.9163 | -106.1478
18.9199 | -106.1476
18.9236 | -106.1494
18.9272 | -106.152
18.9309 | -106.154
18.9346 | -106.1555
18.9382 | -106.1578
18.9419 | -106.1613
18.9455 | -106.163
18.9492 | -106.1636
18.9529 | -106.1613
18.9565 | -106.1579
18.9602 | -106.1564
18.9638 | -106.1538
18.9675 | -106.1524
18.9711 | -106.1522
18.9748 | -106.152
18.9785 | -106.1488
18.9821 | -106.145
18.9858 | -106.1413
18.9894 | -106.138
18.9931 | -106.133
18.9968 | -106.1253
19.0004 | -106.1212
19.0041 | -106.1168
19.0077 | -106.1144
19.0114 | -106.1115
19.015 | -106.1075
19.0187 | -106.105
19.0224 | -106.1027
19.026 | -106.1026
19.0297 | -106.1006
19.0333 | -106.0986
19.037 | -106.098
19.0407 | -106.0963
19.0443 | -106.0945
19.048 | -106.0931
19.0516 | -106.0937
19.0553 | -106.0933
19.0589 | -106.0914
19.0626 | -106.0905
19.0663 | -106.0911
19.0699 | -106.0917
19.0736 | -106.0927
19.0772 | -106.0922
19.0809 | -106.0917
19.0845 | -106.0928
19.0882 | -106.0946
19.0919 | -106.0963
19.0955 | -106.0984
19.0992 | -106.0995
19.1028 | -106.1006
19.1065 | -106.1019
19.1102 | -106.1035
19.1138 | -106.1039
19.1175 | -106.1035
19.1211 | -106.1012
19.1248 | -106.0996
19.1284 | -106.0988
19.1321 | -106.0954
19.1358 | -106.0935
19.1394 | -106.0895
19.1431 | -106.0846
19.1467 | -106.0776
19.1504 | -106.0691
19.1541 | -106.0611
19.1577 | -106.0522
19.1614 | -106.0407
19.165 | -106.0274
19.1687 | -106.0133
19.1723 | -105.9983
19.176 | -105.9806
19.1797 | -105.9629
19.1833 | -105.9471
19.187 | -105.9317
19.1906 | -105.9143
19.1943 | -105.8942
19.198 | -105.8737
19.2016 | -105.8534
19.2053 | -105.8329
19.2089 | -105.812
19.2126 | -105.791
19.2162 | -105.7719
19.2199 | -105.7536
19.2236 | -105.7375
19.2272 | -105.7227
19.2309 | -105.7101
19.2345 | -105.6968
19.2382 | -105.6818
19.2418 | -105.6655
19.2455 | -105.6525
19.2492 | -105.6432
19.2528 | -105.6368
19.2565 | -105.6313
19.2601 | -105.6274
19.2638 | -105.6255
19.2675 | -105.626
19.2711 | -105.6269
19.2748 | -105.628
19.2784 | -105.6322
19.2821 | -105.6377
19.2857 | -105.6445
19.2894 | -105.6539
19.2931 | -105.6636
19.2967 | -105.675
19.3004 | -105.6888
19.304 | -105.7032
19.3077 | -105.7198
19.3114 | -105.738
19.315 | -105.7564
19.3187 | -105.7766
19.3223 | -105.7954
19.326 | -105.8136
19.3296 | -105.8335
19.3333 | -105.8541
19.337 | -105.8739
19.3406 | -105.8925
19.3443 | -105.9077
19.3479 | -105.9256
19.3516 | -105.9446
19.3553 | -105.9608
19.3589 | -105.9732
19.3626 | -105.9824
19.3662 | -105.9933
19.3699 | -106.0011
19.3735 | -106.0078
19.3772 | -106.0149
19.3809 | -106.0203
19.3845 | -106.0245
19.3882 | -106.0268
19.3918 | -106.0251
19.3955 | -106.0247
19.3991 | -106.0221
19.4028 | -106.0155
19.4065 | -106.009
19.4101 | -106.0021
19.4138 | -105.9949
19.4174 | -105.9857
19.4211 | -105.975
19.4248 | -105.9659
19.4284 | -105.9533
19.4321 | -105.9367
19.4357 | -105.9235
19.4394 | -105.9099
19.443 | -105.894
19.4467 | -105.881
19.4504 | -105.8666
19.454 | -105.8521
19.4577 | -105.8371
19.4613 | -105.8225
19.465 | -105.8084
19.4687 | -105.795
19.4723 | -105.7791
19.476 | -105.7648
19.4796 | -105.7513
19.4833 | -105.7387
19.4869 | -105.7284
19.4906 | -105.7174
19.4943 | -105.7065
19.4979 | -105.6931
19.5016 | -105.6804
19.5052 | -105.6676
19.5089 | -105.6549
19.5126 | -105.6445
19.5162 | -105.6372
19.5199 | -105.6277
19.5235 | -105.6175
19.5272 | -105.6066
19.5308 | -105.5952
19.5345 | -105.5832
19.5382 | -105.5705
19.5418 | -105.5599
19.5455 | -105.5501
19.5491 | -105.5398
19.5528 | -105.5313
19.5565 | -105.5218
19.5601 | -105.5117
19.5638 | -105.5021
19.5674 | -105.4938
19.5711 | -105.4859
19.5747 | -105.4794
19.5784 | -105.4702
19.5821 | -105.4625
19.5857 | -105.4567
19.5894 | -105.4492
19.593 | -105.443
19.5967 | -105.4382
19.6003 | -105.4355
19.604 | -105.433
19.6077 | -105.4299
19.6113 | -105.4278
19.615 | -105.4268
19.6186 | -105.4262
19.6223 | -105.4296
19.626 | -105.4323
19.6296 | -105.4342
19.6333 | -105.4367
19.6369 | -105.4399
19.6406 | -105.4472
19.6442 | -105.4526
19.6479 | -105.4585
19.6516 | -105.4648
19.6552 | -105.4711
19.6589 | -105.4792
19.6625 | -105.4866
19.6662 | -105.4948
19.6699 | -105.5042
19.6735 | -105.5122
19.6772 | -105.5225
19.6808 | -105.5341
19.6845 | -105.5461
19.6881 | -105.5568
19.6918 | -105.5672
19.6955 | -105.5764
19.6991 | -105.587
19.7028 | -105.598
19.7064 | -105.6094
19.7101 | -105.6219
19.7138 | -105.631
19.7174 | -105.6419
19.7211 | -105.6535
19.7247 | -105.6623
19.7284 | -105.671
19.732 | -105.6785
19.7357 | -105.6848
19.7394 | -105.6925
19.743 | -105.7003
19.7467 | -105.7061
19.7503 | -105.7125
19.754 | -105.7193
19.7576 | -105.7267
19.7613 | -105.7321
19.765 | -105.7341
19.7686 | -105.7368
19.7723 | -105.7385
19.7759 | -105.7424
19.7796 | -105.7447
19.7833 | -105.7454
19.7869 | -105.7463
19.7906 | -105.7464
19.7942 | -105.7437
19.7979 | -105.7417
19.8015 | -105.7404
19.8052 | -105.7379
19.8089 | -105.7356
19.8125 | -105.7327
19.8162 | -105.7286
19.8198 | -105.7232
19.8235 | -105.7174
19.8272 | -105.7105
19.8308 | -105.7053
19.8345 | -105.6974
19.8381 | -105.6907
19.8418 | -105.6853
19.8454 | -105.679
19.8491 | -105.6704
19.8528 | -105.6627
19.8564 | -105.6557
19.8601 | -105.6505
19.8637 | -105.643
19.8674 | -105.6358
19.8711 | -105.6313
19.8747 | -105.6249
19.8784 | -105.6186
19.882 | -105.6132
19.8857 | -105.6096
19.8893 | -105.6065
19.893 | -105.6033
19.8967 | -105.6008
19.9003 | -105.6002
19.904 | -105.597
19.9076 | -105.5954
19.9113 | -105.5964
19.9149 | -105.5998
19.9186 | -105.6028
19.9223 | -105.6011
19.9259 | -105.6028
19.9296 | -105.606
19.9332 | -105.6104
19.9369 | -105.6135
19.9406 | -105.6185
19.9442 | -105.6251
19.9479 | -105.6318
19.9515 | -105.64
19.9552 | -105.6457
19.9588 | -105.65
19.9625 | -105.6556
19.9662 | -105.6619
19.9698 | -105.6678
19.9735 | -105.6728
19.9771 | -105.6783
19.9808 | -105.6831
19.9845 | -105.6872
19.9881 | -105.6898
19.9918 | -105.6919
19.9954 | -105.6927
19.9991 | -105.689
20.0027 | -105.6864
20.0064 | -105.6841
20.0101 | -105.6814
20.0137 | -105.6774
20.0174 | -105.6721
20.021 | -105.6637
20.0247 | -105.6548
20.0284 | -105.6474
20.032 | -105.6389
20.0357 | -105.6274
20.0393 | -105.6147
20.043 | -105.602
20.0466 | -105.589
20.0503 | -105.5745
20.054 | -105.5584
20.0576 | -105.545
20.0613 | -105.5291
20.0649 | -105.511
20.0686 | -105.494
20.0722 | -105.478
20.0759 | -105.4612
20.0796 | -105.4455
20.0832 | -105.4283
20.0869 | -105.4121
20.0905 | -105.3975
20.0942 | -105.3822
20.0979 | -105.3677
20.1015 | -105.3543
20.1052 | -105.3412
20.1088 | -105.3291
20.1125 | -105.3196
20.1161 | -105.3098
20.1198 | -105.3009
20.1235 | -105.2935
20.1271 | -105.2877
20.1308 | -105.2828
20.1344 | -105.2785
20.1381 | -105.2745
20.1418 | -105.2727
20.1454 | -105.2737
20.1491 | -105.2758
20.1527 | -105.2754
20.1564 | -105.2781
20.16 | -105.283
20.1637 | -105.2873
20.1674 | -105.293
20.171 | -105.2983
20.1747 | -105.3053
20.1783 | -105.314
20.182 | -105.3236
20.1857 | -105.3318
20.1893 | -105.3408
20.193 | -105.3498
20.1966 | -105.3584
20.2003 | -105.3688
20.2039 | -105.3819
20.2076 | -105.3973
20.2113 | -105.4096
20.2149 | -105.4219
20.2186 | -105.4321
20.2222 | -105.4429
20.2259 | -105.4539
20.2295 | -105.4626
20.2332 | -105.4722
20.2369 | -105.481
20.2405 | -105.4922
20.2442 | -105.5048
20.2478 | -105.5168
20.2515 | -105.5269
20.2552 | -105.5357
20.2588 | -105.5465
20.2625 | -105.5587
20.2661 | -105.5718
20.2698 | -105.5827
20.2734 | -105.5939
20.2771 | -105.605
20.2808 | -105.6147
20.2844 | -105.6267
20.2881 | -105.6358
20.2917 | -105.6434
20.2954 | -105.6507
20.2991 | -105.6574
20.3027 | -105.6664
20.3064 | -105.6735
20.31 | -105.6794
20.3137 | -105.6832
20.3173 | -105.6882
20.321 | -105.6923
20.3247 | -105.6957
20.3283 | -105.7002
20.332 | -105.7042
20.3356 | -105.7053
20.3393 | -105.7067
20.343 | -105.7072
20.3466 | -105.7088
20.3503 | -105.7122
20.3539 | -105.7154
20.3576 | -105.7149
20.3612 | -105.7121
20.3649 | -105.708
20.3686 | -105.7045
20.3722 | -105.7002
20.3759 | -105.6929
20.3795 | -105.685
20.3832 | -105.6762
20.3868 | -105.6688
20.3905 | -105.6607
20.3942 | -105.6532
20.3978 | -105.6432
20.4015 | -105.6295
20.4051 | -105.6162
20.4088 | -105.605
20.4125 | -105.5928
20.4161 | -105.5808
20.4198 | -105.568
20.4234 | -105.5558
20.4271 | -105.5444
20.4307 | -105.5317
20.4344 | -105.5188
20.4381 | -105.5088
20.4417 | -105.5001
20.4454 | -105.4887
20.449 | -105.4792
20.4527 | -105.4726
20.4564 | -105.4652
20.46 | -105.4601
20.4637 | -105.453
20.4673 | -105.4459
20.471 | -105.4402
20.4746 | -105.4365
20.4783 | -105.4337
20.482 | -105.4319
20.4856 | -105.4315
20.4893 | -105.4299
20.4929 | -105.4256
20.4966 | -105.4227
20.5003 | -105.4199
20.5039 | -105.4179
20.5076 | -105.4196
20.5112 | -105.4231
20.5149 | -105.4225
20.5185 | -105.422
20.5222 | -105.4224
20.5259 | -105.4236
20.5295 | -105.425
20.5332 | -105.4249
20.5368 | -105.4222
20.5405 | -105.4204
20.5441 | -105.4227
20.5478 | -105.4217
20.5515 | -105.4181
20.5551 | -105.4146
20.5588 | -105.4107
20.5624 | -105.4076
20.5661 | -105.4049
20.5698 | -105.4019
20.5734 | -105.3984
20.5771 | -105.3907
20.5807 | -105.374
20.5844 | -105.3652
20.588 | -105.3549
20.5917 | -105.3443
20.5954 | -105.3329
20.599 | -105.3212
20.6027 | -105.3105
20.6063 | -105.3012
20.61 | -105.2899
20.6137 | -105.2782
20.6173 | -105.266
20.621 | -105.2537
20.6246 | -105.2418
20.6283 | -105.2328
20.6319 | -105.2215
20.6356 | -105.2104
20.6393 | -105.1995
20.6429 | -105.1881
20.6466 | -105.1786
20.6502 | -105.1698
20.6539 | -105.1621
20.6576 | -105.1557
20.6612 | -105.1503
20.6649 | -105.1469
20.6685 | -105.145
20.6722 | -105.1445
20.6758 | -105.1458
20.6795 | -105.1465
20.6832 | -105.1481
20.6868 | -105.1502
20.6905 | -105.1559
20.6941 | -105.1635
20.6978 | -105.1713
20.7015 | -105.1809
20.7051 | -105.1913
20.7088 | -105.2007
20.7124 | -105.2133
20.7161 | -105.228
20.7197 | -105.2429
20.7234 | -105.2568
20.7271 | -105.2709
20.7307 | -105.2864
20.7344 | -105.3038
20.738 | -105.3224
20.7417 | -105.341
20.7453 | -105.3585
20.749 | -105.3773
20.7527 | -105.3957
20.7563 | -105.4174
20.76 | -105.4385
20.7636 | -105.4575
20.7673 | -105.4765
20.771 | -105.4951
20.7746 | -105.5116
20.7783 | -105.5271
20.7819 | -105.5444
20.7856 | -105.5642
20.7892 | -105.5836
20.7929 | -105.6037
20.7966 | -105.6214
20.8002 | -105.6356
20.8039 | -105.6486
20.8075 | -105.6631
20.8112 | -105.6786
20.8149 | -105.6923
20.8185 | -105.704
20.8222 | -105.7122
20.8258 | -105.7209
20.8295 | -105.7321
20.8331 | -105.7414
20.8368 | -105.7467
20.8405 | -105.7535
20.8441 | -105.7606
20.8478 | -105.7641
20.8514 | -105.7691
20.8551 | -105.773
20.8588 | -105.7733
20.8624 | -105.7719
20.8661 | -105.7714
20.8697 | -105.7763
20.8734 | -105.7803
20.877 | -105.7829
20.8807 | -105.7846
20.8844 | -105.7851
20.888 | -105.7846
20.8917 | -105.7825
20.8953 | -105.7808
20.899 | -105.7774
20.9026 | -105.7723
20.9063 | -105.7662
20.91 | -105.7609
20.9136 | -105.7583
20.9173 | -105.755
20.9209 | -105.7524
20.9246 | -105.7509
20.9283 | -105.7486
20.9319 | -105.7439
20.9356 | -105.7414
20.9392 | -105.7399
20.9429 | -105.7362
20.9465 | -105.7337
20.9502 | -105.7321
20.9539 | -105.7307
20.9575 | -105.728
20.9612 | -105.7268
20.9648 | -105.7252
20.9685 | -105.7215
20.9722 | -105.7189
20.9758 | -105.7165
20.9795 | -105.714
20.9831 | -105.712
20.9868 | -105.7089
20.9904 | -105.708
20.9941 | -105.7091
20.9978 | -105.7085
21.0014 | -105.708
21.0051 | -105.7091
21.0087 | -105.71
21.0124 | -105.7085
21.016 | -105.7083
21.0197 | -105.7094
21.0234 | -105.7107
21.027 | -105.7115
21.0307 | -105.71
21.0343 | -105.7088
21.038 | -105.71
21.0417 | -105.7088
21.0453 | -105.7077
21.049 | -105.7088
21.0526 | -105.7089
21.0563 | -105.7078
21.0599 | -105.7077
21.0636 | -105.7097
21.0673 | -105.7115
21.0709 | -105.7133
21.0746 | -105.7137
21.0782 | -105.7109
21.0819 | -105.7085
21.0856 | -105.7069
21.0892 | -105.7028
21.0929 | -105.697
21.0965 | -105.6933
21.1002 | -105.69
21.1038 | -105.6858
21.1075 | -105.6799
21.1112 | -105.6737
21.1148 | -105.6695
21.1185 | -105.6639
21.1221 | -105.6581
21.1258 | -105.6513
21.1295 | -105.644
21.1331 | -105.638
21.1368 | -105.6327
21.1404 | -105.6266
21.1441 | -105.6181
21.1477 | -105.6085
21.1514 | -105.5978
21.1551 | -105.5864
21.1587 | -105.574
21.1624 | -105.5622
21.166 | -105.5484
21.1697 | -105.53
21.1733 | -105.5131
21.177 | -105.4961
21.1807 | -105.4787
21.1843 | -105.4642
21.188 | -105.448
21.1916 | -105.4309
21.1953 | -105.4145
21.199 | -105.3969
21.2026 | -105.3795
21.2063 | -105.364
21.2099 | -105.3488
21.2136 | -105.3322
21.2172 | -105.3142
21.2209 | -105.2958
21.2246 | -105.2771
21.2282 | -105.2602
21.2319 | -105.2454
21.2355 | -105.2282
21.2392 | -105.2116
21.2429 | -105.1987
21.2465 | -105.1851
21.2502 | -105.1716
21.2538 | -105.1581
21.2575 | -105.1451
21.2611 | -105.1292
21.2648 | -105.1175
21.2685 | -105.1085
21.2721 | -105.0982
21.2758 | -105.089
21.2794 | -105.0811
21.2831 | -105.074
21.2868 | -105.0675
21.2904 | -105.0608
21.2941 | -105.0537
21.2977 | -105.0472
21.3014 | -105.0413
21.305 | -105.0369
21.3087 | -105.0347
21.3124 | -105.0336
21.316 | -105.0325
21.3197 | -105.0309
21.3233 | -105.0304
21.327 | -105.0298
21.3307 | -105.0305
21.3343 | -105.0328
21.338 | -105.0337
21.3416 | -105.0369
21.3453 | -105.038
21.3489 | -105.0376
21.3526 | -105.0409
21.3563 | -105.0459
21.3599 | -105.0495
21.3636 | -105.0519
21.3672 | -105.0537
21.3709 | -105.0574
21.3745 | -105.0634
21.3782 | -105.0691
21.3819 | -105.0738
21.3855 | -105.077
21.3892 | -105.0831
21.3928 | -105.0894
21.3965 | -105.0933
21.4002 | -105.0965
21.4038 | -105.0987
21.4075 | -105.1003
21.4111 | -105.1024
21.4148 | -105.1056
21.4184 | -105.1072
21.4221 | -105.1097
21.4258 | -105.1126
21.4294 | -105.1133
21.4331 | -105.1127
21.4367 | -105.1111
21.4404 | -105.1099
21.4441 | -105.1082
21.4477 | -105.1076
21.4514 | -105.1089
21.455 | -105.1093
21.4587 | -105.1073
21.4623 | -105.107
21.466 | -105.1046
21.4697 | -105.102
21.4733 | -105.1027
21.477 | -105.1006
21.4806 | -105.0975
21.4843 | -105.0942
21.488 | -105.0876
21.4916 | -105.0823
21.4953 | -105.0779
21.4989 | -105.0726
21.5026 | -105.0672
21.5062 | -105.0606
21.5099 | -105.0542
21.5136 | -105.0456
21.5172 | -105.0366
21.5209 | -105.0291
21.5245 | -105.0209
21.5282 | -105.0133
21.5318 | -105.0074
21.5355 | -105.0003
21.5392 | -104.9916
21.5428 | -104.9811
21.5465 | -104.9689
21.5501 | -104.9572
21.5538 | -104.9448
21.5575 | -104.9303
21.5611 | -104.9179
21.5648 | -104.9096
21.5684 | -104.8967
21.5721 | -104.8836
21.5757 | -104.87
21.5794 | -104.8553
21.5831 | -104.8401
21.5867 | -104.8246
21.5904 | -104.8071
21.594 | -104.7906
21.5977 | -104.7729
21.6014 | -104.7555
21.605 | -104.7369
21.6087 | -104.7174
21.6123 | -104.6971
21.616 | -104.678
21.6196 | -104.6576
21.6233 | -104.6389
21.627 | -104.6179
21.6306 | -104.5984
21.6343 | -104.5776
21.6379 | -104.5581
21.6416 | -104.5371
21.6453 | -104.517
21.6489 | -104.4961
21.6526 | -104.4779
21.6562 | -104.4603
21.6599 | -104.4416
21.6635 | -104.4226
21.6672 | -104.4055
21.6709 | -104.3902
21.6745 | -104.3772
21.6782 | -104.364
21.6818 | -104.3504
21.6855 | -104.3372
21.6891 | -104.3243
21.6928 | -104.3125
21.6965 | -104.3007
21.7001 | -104.2915
21.7038 | -104.2857
21.7074 | -104.2787
21.7111 | -104.2718
21.7148 | -104.2654
21.7184 | -104.2605
21.7221 | -104.2561
21.7257 | -104.2519
21.7294 | -104.247
21.733 | -104.2432
21.7367 | -104.2424
21.7404 | -104.243
21.744 | -104.2421
21.7477 | -104.2422
21.7513 | -104.2435
21.755 | -104.2453
21.7587 | -104.2482
21.7623 | -104.252
21.766 | -104.2576
21.7696 | -104.262
21.7733 | -104.2674
21.7769 | -104.2737
21.7806 | -104.2794
21.7843 | -104.2872
21.7879 | -104.2959
21.7916 | -104.3045
21.7952 | -104.315
21.7989 | -104.3262
21.8026 | -104.3375
21.8062 | -104.3507
21.8099 | -104.3656
21.8135 | -104.3808
21.8172 | -104.394
21.8208 | -104.409
21.8245 | -104.4246
21.8282 | -104.4401
21.8318 | -104.4571
21.8355 | -104.4733
21.8391 | -104.4888
21.8428 | -104.5047
21.8464 | -104.5234
21.8501 | -104.5428
21.8538 | -104.5602
21.8574 | -104.5771
21.8611 | -104.5955
21.8647 | -104.6155
21.8684 | -104.6339
21.8721 | -104.6505
21.8757 | -104.6671
21.8794 | -104.6823
21.883 | -104.6995
21.8867 | -104.716
21.8903 | -104.7313
21.894 | -104.7442
21.8977 | -104.758
21.9013 | -104.7689
21.905 | -104.7791
21.9086 | -104.7892
21.9123 | -104.8009
21.916 | -104.8073
21.9196 | -104.8114
21.9233 | -104.812
21.9269 | -104.813
21.9306 | -104.8121
21.9342 | -104.8096
21.9379 | -104.8064
21.9416 | -104.8021
21.9452 | -104.7932
21.9489 | -104.7807
21.9525 | -104.7675
21.9562 | -104.7496
21.9599 | -104.7334
21.9635 | -104.7163
21.9672 | -104.6968
21.9708 | -104.6739
21.9745 | -104.6491
21.9781 | -104.6237
21.9818 | -104.5941
21.9855 | -104.5634
21.9891 | -104.5311
21.9928 | -104.4978
21.9964 | -104.4669
22.0001 | -104.4361
22.0038 | -104.4037
22.0074 | -104.3704
22.0111 | -104.3369
22.0147 | -104.3007
22.0184 | -104.2667
22.022 | -104.2339
22.0257 | -104.201
22.0294 | -104.1681
22.033 | -104.135
22.0367 | -104.1022
22.0403 | -104.073
22.044 | -104.0457
22.0476 | -104.0185
22.0513 | -103.9917
22.055 | -103.9689
22.0586 | -103.9471
22.0623 | -103.9269
22.0659 | -103.9089
22.0696 | -103.8912
22.0733 | -103.8771
22.0769 | -103.8657
22.0806 | -103.8595
22.0842 | -103.8561
22.0879 | -103.8528
22.0915 | -103.8521
22.0952 | -103.8532
22.0989 | -103.8563
22.1025 | -103.8616
22.1062 | -103.8703
22.1098 | -103.8823
22.1135 | -103.8952
22.1172 | -103.91
22.1208 | -103.9277
22.1245 | -103.9475
22.1281 | -103.9714
22.1318 | -103.9964
22.1354 | -104.0215
22.1391 | -104.0496
22.1428 | -104.0791
22.1464 | -104.1105
22.1501 | -104.1411
22.1537 | -104.1748
22.1574 | -104.2062
22.161 | -104.2385
22.1647 | -104.2709
22.1684 | -104.3032
22.172 | -104.3358
22.1757 | -104.3669
22.1793 | -104.3978
22.183 | -104.4267
22.1867 | -104.454
22.1903 | -104.4809
22.194 | -104.5089
22.1976 | -104.5358
22.2013 | -104.5624
22.2049 | -104.5869
22.2086 | -104.6116
22.2123 | -104.6347
22.2159 | -104.6575
22.2196 | -104.6799
22.2232 | -104.7012
22.2269 | -104.7187
22.2306 | -104.7386
22.2342 | -104.757
22.2379 | -104.775
22.2415 | -104.7904
22.2452 | -104.8035
22.2488 | -104.8172
22.2525 | -104.8282
22.2562 | -104.8388
22.2598 | -104.8508
22.2635 | -104.8616
22.2671 | -104.8725
22.2708 | -104.8847
22.2745 | -104.8939
22.2781 | -104.8996
22.2818 | -104.9039
22.2854 | -104.9091
22.2891 | -104.916
22.2927 | -104.9216
22.2964 | -104.9285
22.3001 | -104.9346
22.3037 | -104.9402
22.3074 | -104.9461
22.311 | -104.9481
22.3147 | -104.9489
22.3183 | -104.9506
22.322 | -104.9522
22.3257 | -104.9524
22.3293 | -104.9552
22.333 | -104.9556
22.3366 | -104.9534
22.3403 | -104.9528
22.344 | -104.9505
22.3476 | -104.9473
22.3513 | -104.9442
22.3549 | -104.9373
22.3586 | -104.9311
22.3622 | -104.9251
22.3659 | -104.9175
22.3696 | -104.9074
22.3732 | -104.8989
22.3769 | -104.8911
22.3805 | -104.882
22.3842 | -104.8733
22.3879 | -104.8616
22.3915 | -104.8494
22.3952 | -104.8354
22.3988 | -104.8225
22.4025 | -104.8089
22.4061 | -104.7972
22.4098 | -104.7867
22.4135 | -104.7759
22.4171 | -104.7657
22.4208 | -104.7544
22.4244 | -104.7431
22.4281 | -104.7327
22.4318 | -104.7222
22.4354 | -104.7116
22.4391 | -104.702
22.4427 | -104.6923
22.4464 | -104.6843
22.45 | -104.678
22.4537 | -104.6706
22.4574 | -104.6647
22.461 | -104.6596
22.4647 | -104.6566
22.4683 | -104.6521
22.472 | -104.6526
22.4757 | -104.6535
22.4793 | -104.6548
22.483 | -104.6584
22.4866 | -104.6653
22.4903 | -104.6718
22.4939 | -104.6786
22.4976 | -104.6878
22.5013 | -104.6992
22.5049 | -104.7121
22.5086 | -104.726
22.5122 | -104.7377
22.5159 | -104.7503
22.5195 | -104.7654
22.5232 | -104.7835
22.5269 | -104.8012
22.5305 | -104.8183
22.5342 | -104.8364
22.5378 | -104.8543
22.5415 | -104.8722
22.5452 | -104.8917
22.5488 | -104.9101
22.5525 | -104.9286
22.5561 | -104.9468
22.5598 | -104.9645
22.5634 | -104.9844
22.5671 | -105.0021
22.5708 | -105.02
22.5744 | -105.0363
22.5781 | -105.0537
22.5817 | -105.0684
22.5854 | -105.0805
22.5891 | -105.0906
22.5927 | -105.0969
22.5964 | -105.1061
22.6 | -105.1161
22.6037 | -105.1229
22.6073 | -105.1275
22.611 | -105.1294
22.6147 | -105.1305
22.6183 | -105.1299
22.622 | -105.1282
22.6256 | -105.1242
22.6293 | -105.1209
22.633 | -105.1166
22.6366 | -105.1113
22.6403 | -105.1032
22.6439 | -105.0936
22.6476 | -105.0807
22.6512 | -105.0675
22.6549 | -105.0555
22.6586 | -105.0409
22.6622 | -105.0247
22.6659 | -105.009
22.6695 | -104.994
22.6732 | -104.9784
22.6768 | -104.9594
22.6805 | -104.9395
22.6842 | -104.9206
22.6878 | -104.9028
22.6915 | -104.8829
22.6951 | -104.8633
22.6988 | -104.8429
22.7025 | -104.8244
22.7061 | -104.8048
22.7098 | -104.7868
22.7134 | -104.7713
22.7171 | -104.7569
22.7207 | -104.742
22.7244 | -104.73
22.7281 | -104.7175
22.7317 | -104.7056
22.7354 | -104.6967
22.739 | -104.6874
22.7427 | -104.6788
22.7464 | -104.6696
22.75 | -104.6635
22.7537 | -104.6599
22.7573 | -104.6571
22.761 | -104.6555
22.7646 | -104.6548
22.7683 | -104.6563
22.772 | -104.6602
22.7756 | -104.6665
22.7793 | -104.6717
22.7829 | -104.6768
22.7866 | -104.685
22.7903 | -104.6947
22.7939 | -104.7023
22.7976 | -104.7114
22.8012 | -104.7222
22.8049 | -104.732
22.8085 | -104.743
22.8122 | -104.7567
22.8159 | -104.7721
22.8195 | -104.7874
22.8232 | -104.8016
22.8268 | -104.8157
22.8305 | -104.8304
22.8341 | -104.8467
22.8378 | -104.8614
22.8415 | -104.8758
22.8451 | -104.8889
22.8488 | -104.9022
22.8524 | -104.9158
22.8561 | -104.9293
22.8598 | -104.9427
22.8634 | -104.9551
22.8671 | -104.966
22.8707 | -104.9782
22.8744 | -104.9885
22.878 | -104.9976
22.8817 | -105.0057
22.8854 | -105.0124
22.889 | -105.0206
22.8927 | -105.0275
22.8963 | -105.0337
22.9 | -105.0381
22.9037 | -105.0441
22.9073 | -105.0509
22.911 | -105.0571
22.9146 | -105.0635
22.9183 | -105.0668
22.9219 | -105.0669
22.9256 | -105.0675
22.9293 | -105.0683
22.9329 | -105.0694
22.9366 | -105.0697
22.9402 | -105.0695
22.9439 | -105.0693
22.9476 | -105.0714
22.9512 | -105.0724
22.9549 | -105.0726
22.9585 | -105.0715
22.9622 | -105.0723
22.9658 | -105.0741
22.9695 | -105.0757
22.9732 | -105.0765
22.9768 | -105.0767
22.9805 | -105.0743
22.9841 | -105.0721
22.9878 | -105.0689
22.9914 | -105.065
22.9951 | -105.0619
22.9988 | -105.0591
23.0024 | -105.0562
23.0061 | -105.0534
23.0097 | -105.0497
23.0134 | -105.0466
23.0171 | -105.0442
23.0207 | -105.0418
23.0244 | -105.04
23.028 | -105.0346
23.0317 | -105.0275
23.0353 | -105.0222
23.039 | -105.0145
23.0427 | -105.0061
23.0463 | -104.9973
23.05 | -104.9896
23.0536 | -104.9815
23.0573 | -104.9716
23.061 | -104.9621
23.0646 | -104.953
23.0683 | -104.9426
23.0719 | -104.9335
23.0756 | -104.9249
23.0792 | -104.9136
23.0829 | -104.902
23.0866 | -104.8906
23.0902 | -104.8787
23.0939 | -104.8663
23.0975 | -104.8527
23.1012 | -104.8387
23.1049 | -104.826
23.1085 | -104.8137
23.1122 | -104.8032
23.1158 | -104.7936
23.1195 | -104.783
23.1231 | -104.7725
23.1268 | -104.7606
23.1305 | -104.7491
23.1341 | -104.736
23.1378 | -104.7228
23.1414 | -104.7092
23.1451 | -104.6953
23.1487 | -104.6834
23.1524 | -104.6733
23.1561 | -104.6634
23.1597 | -104.6528
23.1634 | -104.6418
23.167 | -104.631
23.1707 | -104.6218
23.1744 | -104.6148
23.178 | -104.6076
23.1817 | -104.5993
23.1853 | -104.5924
23.189 | -104.5869
23.1926 | -104.5814
23.1963 | -104.5745
23.2 | -104.5697
23.2036 | -104.5659
23.2073 | -104.563
23.2109 | -104.5608
23.2146 | -104.5571
23.2183 | -104.5611
23.2219 | -104.559
23.2256 | -104.5581
23.2292 | -104.5572
23.2329 | -104.557
23.2365 | -104.5576
23.2402 | -104.5593
23.2439 | -104.5614
23.2475 | -104.5634
23.2512 | -104.5679
23.2548 | -104.573
23.2585 | -104.5769
23.2622 | -104.5803
23.2658 | -104.5852
23.2695 | -104.5899
23.2731 | -104.5936
23.2768 | -104.5968
23.2804 | -104.6014
23.2841 | -104.6065
23.2878 | -104.6125
23.2914 | -104.618
23.2951 | -104.6227
23.2987 | -104.6265
23.3024 | -104.6318
23.306 | -104.6367
23.3097 | -104.6422
23.3134 | -104.6491
23.317 | -104.6549
23.3207 | -104.6584
23.3243 | -104.6631
23.328 | -104.6696
23.3317 | -104.6734
23.3353 | -104.6781
23.339 | -104.6843
23.3426 | -104.689
23.3463 | -104.6911
23.3499 | -104.6928
23.3536 | -104.6941
23.3573 | -104.699
23.3609 | -104.7042
23.3646 | -104.7069
23.3682 | -104.7106
23.3719 | -104.7144
23.3756 | -104.7188
23.3792 | -104.7215
23.3829 | -104.7226
23.3865 | -104.7261
23.3902 | -104.7285
23.3938 | -104.7316
23.3975 | -104.7348
23.4012 | -104.7374
23.4048 | -104.738
23.4085 | -104.7405
23.4121 | -104.744
23.4158 | -104.7478
23.4195 | -104.75
23.4231 | -104.7518
23.4268 | -104.7545
23.4304 | -104.7567
23.4341 | -104.7578
23.4377 | -104.7596
23.4414 | -104.7603
23.4451 | -104.7598
23.4487 | -104.7603
23.4524 | -104.759
23.456 | -104.7574
23.4597 | -104.7575
23.4633 | -104.7582
23.467 | -104.759
23.4707 | -104.7614
23.4743 | -104.7641
23.478 | -104.7659
23.4816 | -104.766
23.4853 | -104.767
23.489 | -104.7682
23.4926 | -104.7717
23.4963 | -104.7752
23.4999 | -104.7779
23.5036 | -104.78
23.5072 | -104.7837
23.5109 | -104.7868
23.5146 | -104.7886
23.5182 | -104.7904
23.5219 | -104.7932
23.5255 | -104.7968
23.5292 | -104.7996
23.5329 | -104.8029
23.5365 | -104.806
23.5402 | -104.8073
23.5438 | -104.8072
23.5475 | -104.8057
23.5511 | -104.8042
23.5548 | -104.804
23.5585 | -104.8055
23.5621 | -104.8068
23.5658 | -104.8064
23.5694 | -104.8054
23.5731 | -104.8038
23.5768 | -104.8039
23.5804 | -104.8032
23.5841 | -104.801
23.5877 | -104.7989
23.5914 | -104.7974
23.595 | -104.7961
23.5987 | -104.7937
23.6024 | -104.7909
23.606 | -104.7896
23.6097 | -104.7881
23.6133 | -104.7848
23.617 | -104.7823
23.6207 | -104.7797
23.6243 | -104.7768
23.628 | -104.7765
23.6316 | -104.7769
23.6353 | -104.7772
23.6389 | -104.7776
23.6426 | -104.7799
23.6463 | -104.7803
23.6499 | -104.7827
23.6536 | -104.7848
23.6572 | -104.7842
23.6609 | -104.7842
23.6645 | -104.7867
23.6682 | -104.7871
23.6719 | -104.7874
23.6755 | -104.79
23.6792 | -104.7938
23.6828 | -104.7974
23.6865 | -104.8031
23.6902 | -104.8081
23.6938 | -104.8139
23.6975 | -104.8194
23.7011 | -104.824
23.7048 | -104.8301
23.7084 | -104.8357
23.7121 | -104.8403
23.7158 | -104.8469
23.7194 | -104.854
23.7231 | -104.861
23.7267 | -104.8678
23.7304 | -104.8756
23.7341 | -104.8835
23.7377 | -104.8916
23.7414 | -104.8991
23.745 | -104.9067
23.7487 | -104.9153
23.7523 | -104.9222
23.756 | -104.9303
23.7597 | -104.9394
23.7633 | -104.948
23.767 | -104.957
23.7706 | -104.9678
23.7743 | -104.9779
23.778 | -104.9869
23.7816 | -104.9969
23.7853 | -105.0058
23.7889 | -105.0177
23.7926 | -105.0305
23.7962 | -105.0399
23.7999 | -105.0503
23.8036 | -105.0604
23.8072 | -105.0684
23.8109 | -105.0765
23.8145 | -105.0845
23.8182 | -105.0935
23.8218 | -105.1012
23.8255 | -105.1082
23.8292 | -105.1151
23.8328 | -105.1218
23.8365 | -105.1297
23.8401 | -105.1368
23.8438 | -105.1422
23.8475 | -105.1488
23.8511 | -105.1567
23.8548 | -105.164
23.8584 | -105.1699
23.8621 | -105.1765
23.8657 | -105.1842
23.8694 | -105.1896
23.8731 | -105.1963
23.8767 | -105.2041
23.8804 | -105.2146
23.884 | -105.2257
23.8877 | -105.2342
23.8914 | -105.242
23.895 | -105.2471
23.8987 | -105.2527
23.9023 | -105.2589
23.906 | -105.2659
23.9096 | -105.2724
23.9133 | -105.2789
23.917 | -105.2863
23.9206 | -105.2937
23.9243 | -105.3022
23.9279 | -105.3118
23.9316 | -105.3202
23.9353 | -105.3275
23.9389 | -105.3355
23.9426 | -105.344
23.9462 | -105.3522
23.9499 | -105.3613
23.9535 | -105.3686
23.9572 | -105.3758
23.9609 | -105.3829
23.9645 | -105.387
23.9682 | -105.3907
23.9718 | -105.3934
23.9755 | -105.3981
23.9791 | -105.4027
23.9828 | -105.4075
23.9865 | -105.4098
23.9901 | -105.4132
23.9938 | -105.4169
23.9974 | -105.4187
24.0011 | -105.4191
24.0048 | -105.4203
24.0084 | -105.4207
24.0121 | -105.4207
24.0157 | -105.4197
24.0194 | -105.4174
24.023 | -105.4113
24.0267 | -105.4043
24.0304 | -105.3993
24.034 | -105.3932
24.0377 | -105.3872
24.0413 | -105.3801
24.045 | -105.3731
24.0487 | -105.3679
24.0523 | -105.3628
24.056 | -105.3541
24.0596 | -105.346
24.0633 | -105.3394
24.0669 | -105.3312
24.0706 | -105.3217
24.0743 | -105.3137
24.0779 | -105.3054
24.0816 | -105.2961
24.0852 | -105.287
24.0889 | -105.2805
24.0926 | -105.2731
24.0962 | -105.2667
24.0999 | -105.2577
24.1035 | -105.2485
24.1072 | -105.2401
24.1108 | -105.2317
24.1145 | -105.2249
24.1182 | -105.2191
24.1218 | -105.2139
24.1255 | -105.2104
24.1291 | -105.2055
24.1328 | -105.2008
24.1364 | -105.197
24.1401 | -105.1911
24.1438 | -105.1872
24.1474 | -105.1862
24.1511 | -105.187
24.1547 | -105.1897
24.1584 | -105.1927
24.1621 | -105.1958
24.1657 | -105.2003
24.1694 | -105.204
24.173 | -105.2086
24.1767 | -105.2146
24.1803 | -105.222
24.184 | -105.2312
24.1877 | -105.2418
24.1913 | -105.2522
24.195 | -105.2622
24.1986 | -105.2744
24.2023 | -105.2862
24.206 | -105.297
24.2096 | -105.3105
24.2133 | -105.3218
24.2169 | -105.334
24.2206 | -105.3469
24.2242 | -105.3598
24.2279 | -105.374
24.2316 | -105.3874
24.2352 | -105.4027
24.2389 | -105.4165
24.2425 | -105.4301
24.2462 | -105.444
24.2499 | -105.4582
24.2535 | -105.4703
24.2572 | -105.4811
24.2608 | -105.4924
24.2645 | -105.5038
24.2681 | -105.5146
24.2718 | -105.5242
24.2755 | -105.5159
24.2791 | -105.5119
24.2828 | -105.5069
24.2864 | -105.4953
24.2901 | -105.4741
24.2937 | -105.456
24.2974 | -105.4412
24.3011 | -105.4272
24.3047 | -105.4143
24.3084 | -105.4023
24.312 | -105.3879
24.3157 | -105.3712
24.3194 | -105.3518
24.323 | -105.3298
24.3267 | -105.3069
24.3303 | -105.2835
24.334 | -105.2574
24.3376 | -105.2286
24.3413 | -105.1976
24.345 | -105.1666
24.3486 | -105.1327
24.3523 | -105.0965
24.3559 | -105.0594
24.3596 | -105.0218
24.3633 | -104.9835
24.3669 | -104.9447
24.3706 | -104.9053
24.3742 | -104.8621
24.3779 | -104.8196
24.3815 | -104.7765
24.3852 | -104.7332
24.3889 | -104.6909
24.3925 | -104.648
24.3962 | -104.6045
24.3998 | -104.5642
24.4035 | -104.5246
24.4072 | -104.4854
24.4108 | -104.4474
24.4145 | -104.4111
24.4181 | -104.376
24.4218 | -104.3426
24.4254 | -104.3111
24.4291 | -104.2824
24.4328 | -104.2554
24.4364 | -104.23
24.4401 | -104.2069
24.4437 | -104.1881
24.4474 | -104.1727
24.451 | -104.1604
24.4547 | -104.1494
24.4584 | -104.1402
24.462 | -104.1337
24.4657 | -104.132
24.4693 | -104.1338
24.473 | -104.138
24.4767 | -104.1448
24.4803 | -104.1545
24.484 | -104.1674
24.4876 | -104.1814
24.4913 | -104.1961
24.4949 | -104.2134
24.4986 | -104.2333
24.5023 | -104.2546
24.5059 | -104.2768
24.5096 | -104.3023
24.5132 | -104.328
24.5169 | -104.3548
24.5206 | -104.3819
24.5242 | -104.4094
24.5279 | -104.4377
24.5315 | -104.4673
24.5352 | -104.4962
24.5388 | -104.5265
24.5425 | -104.5565
24.5462 | -104.586
24.5498 | -104.615
24.5535 | -104.6439
24.5571 | -104.6721
24.5608 | -104.6988
24.5645 | -104.723
24.5681 | -104.7395
24.5718 | -104.739
24.5754 | -104.7184
24.5791 | -104.6828
24.5827 | -104.6248
24.5864 | -104.5471
24.5901 | -104.4205
24.5937 | -104.2897
24.5974 | -104.1861
24.601 | -104.1058
24.6047 | -104.0302
24.6084 | -103.9466
24.612 | -103.8556
24.6157 | -103.755
24.6193 | -103.6464
24.623 | -103.5312
24.6266 | -103.4148
24.6303 | -103.2959
24.634 | -103.1762
24.6376 | -103.0555
24.6413 | -102.9265
24.6449 | -102.7909
24.6486 | -102.6507
24.6522 | -102.5076
24.6559 | -102.3618
24.6596 | -102.2166
24.6632 | -102.0721
24.6669 | -101.9277
24.6705 | -101.7834
24.6742 | -101.6405
24.6779 | -101.5006
24.6815 | -101.3634
24.6852 | -101.2295
24.6888 | -101.099
24.6925 | -100.9723
24.6961 | -100.8499
24.6998 | -100.7323
24.7035 | -100.6195
24.7071 | -100.512
24.7108 | -100.4101
24.7144 | -100.3141
24.7181 | -100.2241
24.7218 | -100.1404
24.7254 | -100.0627
24.7291 | -99.9917
24.7327 | -99.9275
24.7364 | -99.8697
24.74 | -99.8191
24.7437 | -99.7763
24.7474 | -99.7404
24.751 | -99.7118
24.7547 | -99.6903
24.7583 | -99.6768
24.762 | -99.6708
24.7657 | -99.6719
24.7693 | -99.6809
24.773 | -99.6973
24.7766 | -99.7209
24.7803 | -99.752
24.7839 | -99.7905
24.7876 | -99.8363
24.7913 | -99.8891
24.7949 | -99.949
24.7986 | -100.0155
24.8022 | -100.1324
24.8059 | -100.2084
24.8095 | -100.2902
24.8132 | -100.3783
24.8169 | -100.4716
24.8205 | -100.57
24.8242 | -100.6735
24.8278 | -100.7816
24.8315 | -100.8949
24.8352 | -101.0134
24.8388 | -101.136
24.8425 | -101.2621
24.8461 | -101.3906
24.8498 | -101.5208
24.8534 | -101.6547
24.8571 | -101.7917
24.8608 | -101.9286
24.8644 | -102.0671
24.8681 | -102.2059
24.8717 | -102.3447
24.8754 | -102.4834
24.8791 | -102.6213
24.8827 | -102.7569
24.8864 | -102.8915
24.89 | -103.0227
24.8937 | -103.1512
24.8973 | -103.277
24.901 | -103.3984
24.9047 | -103.5155
24.9083 | -103.6286
24.912 | -103.7365
24.9156 | -103.8411
24.9193 | -103.9464
24.923 | -104.0571
24.9266 | -104.1733
24.9303 | -104.2887
24.9339 | -104.409
24.9376 | -104.5299
24.9412 | -104.6767
24.9449 | -104.8043
24.9486 | -104.8818
24.9522 | -104.9178
24.9559 | -104.9363
24.9595 | -104.9527
24.9632 | -104.9684
24.9668 | -104.9881
24.9705 | -105.0093
24.9742 | -105.0304
24.9778 | -105.042
24.9815 | -105.0469
24.9851 | -105.0472
24.9888 | -105.0484
24.9925 | -105.0504
24.9961 | -105.0518
24.9998 | -105.0535
25.0034 | -105.0539
25.0071 | -105.0563
25.0107 | -105.0568
25.0144 | -105.0582
25.0181 | -105.0606
25.0217 | -105.0643
25.0254 | -105.0665
25.029 | -105.0675
25.0327 | -105.0698
25.0364 | -105.0716
25.04 | -105.0728
25.0437 | -105.0743
25.0473 | -105.0765
25.051 | -105.078
25.0546 | -105.0787
25.0583 | -105.0805
25.062 | -105.0833
25.0656 | -105.0864
25.0693 | -105.089
25.0729 | -105.0917
25.0766 | -105.0939
25.0803 | -105.0968
25.0839 | -105.0982
25.0876 | -105.1008
25.0912 | -105.1036
25.0949 | -105.1046
25.0985 | -105.1062
25.1022 | -105.1085
25.1059 | -105.1102
25.1095 | -105.1123
25.1132 | -105.1136
25.1168 | -105.1155
25.1205 | -105.1162
25.1241 | -105.1188
25.1278 | -105.1208
25.1315 | -105.121
25.1351 | -105.1213
25.1388 | -105.1233
25.1424 | -105.1238
25.1461 | -105.1237
25.1498 | -105.1235
25.1534 | -105.1232
25.1571 | -105.1251
25.1607 | -105.1275
25.1644 | -105.1306
25.168 | -105.1319
25.1717 | -105.1308
25.1754 | -105.1313
25.179 | -105.1346
25.1827 | -105.138
25.1863 | -105.1387
25.19 | -105.1381
25.1937 | -105.1373
25.1973 | -105.1363
25.201 | -105.138
25.2046 | -105.1403
25.2083 | -105.1413
25.2119 | -105.1412
25.2156 | -105.1447
25.2193 | -105.1479
25.2229 | -105.1503
25.2266 | -105.1519
25.2302 | -105.1532
25.2339 | -105.1533
25.2376 | -105.1552
25.2412 | -105.1571
25.2449 | -105.1598
25.2485 | -105.1618
25.2522 | -105.1639
25.2558 | -105.167
25.2595 | -105.1687
25.2632 | -105.1705
25.2668 | -105.1737
25.2705 | -105.1776
25.2741 | -105.1791
25.2778 | -105.1807
25.2814 | -105.1839
25.2851 | -105.1873
25.2888 | -105.1892
25.2924 | -105.1904
25.2961 | -105.1909
25.2997 | -105.1897
25.3034 | -105.1879
25.3071 | -105.186
25.3107 | -105.1848
25.3144 | -105.1824
25.318 | -105.1797
25.3217 | -105.1765
25.3253 | -105.1747
25.329 | -105.1649
25.3327 | -105.1636
25.3363 | -105.1638
25.34 | -105.1611
25.3436 | -105.157
25.3473 | -105.1504
25.351 | -105.1438
25.3546 | -105.1369
25.3583 | -105.1288
25.3619 | -105.1213
25.3656 | -105.1139
25.3692 | -105.1062
25.3729 | -105.0973
25.3766 | -105.0866
25.3802 | -105.0764
25.3839 | -105.0675
25.3875 | -105.0575
25.3912 | -105.0472
25.3949 | -105.0368
25.3985 | -105.0266
25.4022 | -105.0166
25.4058 | -105.0061
25.4095 | -104.9951
25.4131 | -104.9829
25.4168 | -104.9702
25.4205 | -104.9563
25.4241 | -104.9439
25.4278 | -104.9323
25.4314 | -104.9199
25.4351 | -104.9069
25.4387 | -104.8949
25.4424 | -104.8845
25.4461 | -104.8731
25.4497 | -104.8617
25.4534 | -104.8514
25.457 | -104.8409
25.4607 | -104.8302
25.4644 | -104.8223
25.468 | -104.8123
25.4717 | -104.8016
25.4753 | -104.7932
25.479 | -104.7852
25.4826 | -104.7752
25.4863 | -104.7667
25.49 | -104.7587
25.4936 | -104.7512
25.4973 | -104.7438
25.5009 | -104.7372
25.5046 | -104.7311
25.5083 | -104.7248
25.5119 | -104.7195
25.5156 | -104.7137
25.5192 | -104.7081
25.5229 | -104.7026
25.5265 | -104.6978
25.5302 | -104.6958
25.5339 | -104.6934
25.5375 | -104.6886
25.5412 | -104.6841
25.5448 | -104.6813
25.5485 | -104.6786
25.5522 | -104.6767
25.5558 | -104.6757
25.5595 | -104.6745
25.5631 | -104.6754
25.5668 | -104.6757
25.5704 | -104.6756
25.5741 | -104.6744
25.5778 | -104.6749
25.5814 | -104.6761
25.5851 | -104.6783
25.5887 | -104.6815
25.5924 | -104.6841
25.596 | -104.6858
25.5997 | -104.6894
25.6034 | -104.693
25.607 | -104.6955
25.6107 | -104.6977
25.6143 | -104.7016
25.618 | -104.706
25.6217 | -104.7119
25.6253 | -104.7155
25.629 | -104.7189
25.6326 | -104.7224
25.6363 | -104.7272
25.6399 | -104.7313
25.6436 | -104.7342
25.6473 | -104.7383
25.6509 | -104.7427
25.6546 | -104.7466
25.6582 | -104.7529
25.6619 | -104.7585
25.6656 | -104.7649
25.6692 | -104.7699
25.6729 | -104.7766
25.6765 | -104.7823
25.6802 | -104.7875
25.6838 | -104.7912
25.6875 | -104.7946
25.6912 | -104.7981
25.6948 | -104.7998
25.6985 | -104.7994
25.7021 | -104.8008
25.7058 | -104.8025
25.7095 | -104.8023
25.7131 | -104.8011
25.7168 | -104.8024
25.7204 | -104.8045
25.7241 | -104.8059
25.7277 | -104.8059
25.7314 | -104.8034
25.7351 | -104.8007
25.7387 | -104.801
25.7424 | -104.8002
25.746 | -104.7985
25.7497 | -104.7969
25.7533 | -104.7958
25.757 | -104.7942
25.7607 | -104.7919
25.7643 | -104.7889
25.768 | -104.7867
25.7716 | -104.7848
25.7753 | -104.7835
25.779 | -104.7832
25.7826 | -104.782
25.7863 | -104.7778
25.7899 | -104.7747
25.7936 | -104.7744
25.7972 | -104.7751
25.8009 | -104.7735
25.8046 | -104.7701
25.8082 | -104.7666
25.8119 | -104.7645
25.8155 | -104.7632
25.8192 | -104.7611
25.8229 | -104.7597
25.8265 | -104.7573
25.8302 | -104.7565
25.8338 | -104.7563
25.8375 | -104.7517
25.8411 | -104.7494
25.8448 | -104.7496
25.8485 | -104.7487
25.8521 | -104.7491
25.8558 | -104.749
25.8594 | -104.7489
25.8631 | -104.7472
25.8668 | -104.7446
25.8704 | -104.7435
25.8741 | -104.744
25.8777 | -104.7433
25.8814 | -104.7439
25.885 | -104.7443
25.8887 | -104.745
25.8924 | -104.7457
25.896 | -104.7454
25.8997 | -104.7464
25.9033 | -104.7481
25.907 | -104.7498
25.9107 | -104.7518
25.9143 | -104.7533
25.918 | -104.755
25.9216 | -104.7566
25.9253 | -104.7578
25.9289 | -104.7608
25.9326 | -104.7637
25.9363 | -104.7665
25.9399 | -104.7689
25.9436 | -104.7711
25.9472 | -104.7743
25.9509 | -104.778
25.9545 | -104.7794
25.9582 | -104.7833
25.9619 | -104.7866
25.9655 | -104.7881
25.9692 | -104.7921
25.9728 | -104.7981
25.9765 | -104.8043
25.9802 | -104.8095
25.9838 | -104.8155
25.9875 | -104.8195
25.9911 | -104.8244
25.9948 | -104.8282
25.9984 | -104.8323
26.0021 | -104.8364
26.0058 | -104.84
26.0094 | -104.8437
26.0131 | -104.8463
26.0167 | -104.8498
26.0204 | -104.8538
26.0241 | -104.8562
26.0277 | -104.8573
26.0314 | -104.8572
26.035 | -104.8586
26.0387 | -104.8584
26.0423 | -104.8593
26.046 | -104.8603
26.0497 | -104.8604
26.0533 | -104.8619
26.057 | -104.8624
26.0606 | -104.8632
26.0643 | -104.8627
26.0679 | -104.8613
26.0716 | -104.8605
26.0753 | -104.8608
26.0789 | -104.8611
26.0826 | -104.859
26.0862 | -104.8547
26.0899 | -104.8499
26.0936 | -104.8467
26.0972 | -104.8462
26.1009 | -104.8439
26.1045 | -104.8393
26.1082 | -104.8344
26.1118 | -104.8289
26.1155 | -104.8239
26.1192 | -104.8183
26.1228 | -104.8142
26.1265 | -104.8089
26.1301 | -104.8031
26.1338 | -104.7968
26.1375 | -104.7906
26.1411 | -104.7847
26.1448 | -104.7765
26.1484 | -104.7703
26.1521 | -104.7659
26.1557 | -104.7593
26.1594 | -104.7516
26.1631 | -104.7436
26.1667 | -104.7362
26.1704 | -104.73
26.174 | -104.724
26.1777 | -104.7173
26.1814 | -104.7094
26.185 | -104.7021
26.1887 | -104.6961
26.1923 | -104.6902
26.196 | -104.6841
26.1996 | -104.68
26.2033 | -104.6782
26.207 | -104.6742
26.2106 | -104.6704
26.2143 | -104.6671
26.2179 | -104.6634
26.2216 | -104.6605
26.2252 | -104.6594
26.2289 | -104.659
26.2326 | -104.6595
26.2362 | -104.6613
26.2399 | -104.6608
26.2435 | -104.6592
26.2472 | -104.6597
26.2509 | -104.66
26.2545 | -104.6617
26.2582 | -104.664
26.2618 | -104.6651
26.2655 | -104.6678
26.2691 | -104.671
26.2728 | -104.6759
26.2765 | -104.6815
26.2801 | -104.6863
26.2838 | -104.6926
26.2874 | -104.6996
26.2911 | -104.7074
26.2948 | -104.7135
26.2984 | -104.7176
26.3021 | -104.7233
26.3057 | -104.7324
26.3094 | -104.742
26.313 | -104.7494
26.3167 | -104.7575
26.3204 | -104.7662
26.324 | -104.7742
26.3277 | -104.7827
26.3313 | -104.7919
26.335 | -104.801
26.3387 | -104.8083
26.3423 | -104.8138
26.346 | -104.8205
26.3496 | -104.8264
26.3533 | -104.8297
26.3569 | -104.8355
26.3606 | -104.8426
26.3643 | -104.8495
26.3679 | -104.8558
26.3716 | -104.861
26.3752 | -104.8647
26.3789 | -104.867
26.3826 | -104.8701
26.3862 | -104.8671
26.3899 | -104.8699
26.3935 | -104.8715
26.3972 | -104.8717
26.4008 | -104.8726
26.4045 | -104.8735
26.4082 | -104.8726
26.4118 | -104.8722
26.4155 | -104.8713
26.4191 | -104.8697
26.4228 | -104.8667
26.4264 | -104.8627
26.4301 | -104.8596
26.4338 | -104.8559
26.4374 | -104.8537
26.4411 | -104.8525
26.4447 | -104.848
26.4484 | -104.8443
26.4521 | -104.8415
26.4557 | -104.8385
26.4594 | -104.8363
26.463 | -104.8346
26.4667 | -104.8302
26.4703 | -104.8266
26.474 | -104.8242
26.4777 | -104.8225
26.4813 | -104.8215
26.485 | -104.8207
26.4886 | -104.8193
26.4923 | -104.8174
26.496 | -104.8155
26.4996 | -104.8152
26.5033 | -104.8139
26.5069 | -104.8128
26.5106 | -104.8119
26.5142 | -104.8126
26.5179 | -104.8148
26.5216 | -104.8172
26.5252 | -104.82
26.5289 | -104.8221
26.5325 | -104.8235
26.5362 | -104.826
26.5399 | -104.8293
26.5435 | -104.8311
26.5472 | -104.834
26.5508 | -104.8365
26.5545 | -104.8403
26.5581 | -104.8451
26.5618 | -104.8484
26.5655 | -104.8522
26.5691 | -104.8556
26.5728 | -104.8593
26.5764 | -104.8634
26.5801 | -104.8685
26.5837 | -104.873
26.5874 | -104.8778
26.5911 | -104.8837
26.5947 | -104.8896
26.5984 | -104.8936
26.602 | -104.8966
26.6057 | -104.9004
26.6094 | -104.9038
26.613 | -104.9083
26.6167 | -104.9127
26.6203 | -104.9171
26.624 | -104.9213
26.6276 | -104.9251
26.6313 | -104.9287
26.635 | -104.9334
26.6386 | -104.9365
26.6423 | -104.9383
26.6459 | -104.9406
26.6496 | -104.9436
26.6533 | -104.9466
26.6569 | -104.9497
26.6606 | -104.9507
26.6642 | -104.9513
26.6679 | -104.9539
26.6715 | -104.9565
26.6752 | -104.9581
26.6789 | -104.9583
26.6825 | -104.9586
26.6862 | -104.96
26.6898 | -104.9615
26.6935 | -104.9609
26.6972 | -104.9613
26.7008 | -104.9612
26.7045 | -104.9606
26.7081 | -104.9602
26.7118 | -104.9603
26.7154 | -104.9623
26.7191 | -104.9616
26.7228 | -104.9619
26.7264 | -104.9652
26.7301 | -104.9653
26.7337 | -104.9636
26.7374 | -104.9618
26.741 | -104.9619
26.7447 | -104.9621
26.7484 | -104.9636
26.752 | -104.9655
26.7557 | -104.9662
26.7593 | -104.9659
26.763 | -104.9662
26.7667 | -104.9646
26.7703 | -104.9637
26.774 | -104.9643
26.7776 | -104.9642
26.7813 | -104.9636
26.7849 | -104.9637
26.7886 | -104.9646
26.7923 | -104.9651
26.7959 | -104.9644
26.7996 | -104.9654
26.8032 | -104.9663
26.8069 | -104.9659
26.8106 | -104.9647
26.8142 | -104.964
26.8179 | -104.9638
26.8215 | -104.9642
26.8252 | -104.9642
26.8288 | -104.9646
26.8325 | -104.965
26.8362 | -104.9668
26.8398 | -104.9686
26.8435 | -104.972
26.8471 | -104.9752
26.8508 | -104.9775
26.8545 | -104.9785
26.8581 | -104.9799
26.8618 | -104.9813
26.8654 | -104.9812
26.8691 | -104.9823
26.8727 | -104.9846
26.8764 | -104.987
26.8801 | -104.991
26.8837 | -104.9947
26.8874 | -104.9983
26.891 | -104.9996
26.8947 | -105.0002
26.8983 | -105.0006
26.902 | -105.0016
26.9057 | -105.0053
26.9093 | -105.0091
26.913 | -105.0137
26.9166 | -105.0165
26.9203 | -105.0186
26.924 | -105.0231
26.9276 | -105.0269
26.9313 | -105.03
26.9349 | -105.0323
26.9386 | -105.0325
26.9422 | -105.0329
26.9459 | -105.034
26.9496 | -105.0362
26.9532 | -105.0384
26.9569 | -105.0401
26.9605 | -105.0406
26.9642 | -105.0425
26.9679 | -105.0443
26.9715 | -105.0458
26.9752 | -105.0481
26.9788 | -105.0518
26.9825 | -105.057
26.9861 | -105.0614
26.9898 | -105.0646
26.9935 | -105.0676
26.9971 | -105.0725
27.0008 | -105.077
27.0044 | -105.0793
27.0081 | -105.083
27.0118 | -105.0854
27.0154 | -105.0846
27.0191 | -105.0837
27.0227 | -105.0862
27.0264 | -105.0869
27.03 | -105.0868
27.0337 | -105.0879
27.0374 | -105.0876
27.041 | -105.087
27.0447 | -105.0873
27.0483 | -105.0867
27.052 | -105.0835
27.0556 | -105.0803
27.0593 | -105.0739
27.063 | -105.0686
27.0666 | -105.0624
27.0703 | -105.0553
27.0739 | -105.0495
27.0776 | -105.0458
27.0813 | -105.0392
27.0849 | -105.0311
27.0886 | -105.0222
27.0922 | -105.0125
27.0959 | -105.0029
27.0995 | -104.9912
27.1032 | -104.9779
27.1069 | -104.9642
27.1105 | -104.9524
27.1142 | -104.9395
27.1178 | -104.9279
27.1215 | -104.9144
27.1252 | -104.8995
27.1288 | -104.8841
27.1325 | -104.8693
27.1361 | -104.8543
27.1398 | -104.8385
27.1434 | -104.8222
27.1471 | -104.8079
27.1508 | -104.7929
27.1544 | -104.7784
27.1581 | -104.7651
27.1617 | -104.751
27.1654 | -104.7377
27.1691 | -104.7245
27.1727 | -104.7113
27.1764 | -104.6998
27.18 | -104.6881
27.1837 | -104.6754
27.1873 | -104.6635
27.191 | -104.653
27.1947 | -104.6431
27.1983 | -104.632
27.202 | -104.622
27.2056 | -104.6117
27.2093 | -104.6008
27.2129 | -104.5912
27.2166 | -104.5821
27.2203 | -104.5747
27.2239 | -104.5681
27.2276 | -104.5623
27.2312 | -104.5574
27.2349 | -104.5528
27.2386 | -104.5501
27.2422 | -104.5492
27.2459 | -104.5473
27.2495 | -104.5452
27.2532 | -104.5443
27.2568 | -104.5452
27.2605 | -104.5454
27.2642 | -104.5438
27.2678 | -104.5433
27.2715 | -104.5452
27.2751 | -104.5491
27.2788 | -104.553
27.2825 | -104.5558
27.2861 | -104.5581
27.2898 | -104.5611
27.2934 | -104.5651
27.2971 | -104.5698
27.3007 | -104.5765
27.3044 | -104.5835
27.3081 | -104.5909
27.3117 | -104.5979
27.3154 | -104.6054
27.319 | -104.6129
27.3227 | -104.6214
27.3264 | -104.6308
27.33 | -104.6403
27.3337 | -104.6491
27.3373 | -104.6569
27.341 | -104.6656
27.3446 | -104.674
27.3483 | -104.6819
27.352 | -104.6891
27.3556 | -104.6972
27.3593 | -104.7055
27.3629 | -104.7136
27.3666 | -104.7194
27.3702 | -104.7252
27.3739 | -104.7321
27.3776 | -104.7403
27.3812 | -104.7468
27.3849 | -104.7513
27.3885 | -104.7576
27.3922 | -104.7628
27.3959 | -104.7688
27.3995 | -104.7769
27.4032 | -104.7837
27.4068 | -104.7866
27.4105 | -104.7912
27.4141 | -104.7963
27.4178 | -104.7991
27.4215 | -104.8022
27.4251 | -104.8056
27.4288 | -104.8081
27.4324 | -104.8111
27.4361 | -104.8148
27.4398 | -104.8133
27.4434 | -104.8168
27.4471 | -104.8213
27.4507 | -104.8211
27.4544 | -104.8216
27.458 | -104.8231
27.4617 | -104.8229
27.4654 | -104.8223
27.469 | -104.8225
27.4727 | -104.8231
27.4763 | -104.8217
27.48 | -104.8194
27.4837 | -104.8173
27.4873 | -104.818
27.491 | -104.8183
27.4946 | -104.8185
27.4983 | -104.8179
27.5019 | -104.8181
27.5056 | -104.8158
27.5093 | -104.8138
27.5129 | -104.813
27.5166 | -104.8112
27.5202 | -104.808
27.5239 | -104.8031
27.5276 | -104.8001
27.5312 | -104.798
27.5349 | -104.7951
27.5385 | -104.7907
27.5422 | -104.7856
27.5458 | -104.7802
27.5495 | -104.7759
27.5532 | -104.7715
27.5568 | -104.7669
27.5605 | -104.7616
27.5641 | -104.755
27.5678 | -104.7502
27.5714 | -104.7464
27.5751 | -104.7401
27.5788 | -104.7344
27.5824 | -104.7286
27.5861 | -104.7222
27.5897 | -104.7164
27.5934 | -104.7117
27.5971 | -104.7072
27.6007 | -104.7026
27.6044 | -104.6984
27.608 | -104.6929
27.6117 | -104.6865
27.6153 | -104.68
27.619 | -104.6732
27.6227 | -104.668
27.6263 | -104.6614
27.63 | -104.6549
27.6336 | -104.6521
27.6373 | -104.6489
27.641 | -104.6426
27.6446 | -104.6377
27.6483 | -104.6341
27.6519 | -104.6334
27.6556 | -104.6333
27.6592 | -104.6314
27.6629 | -104.6287
27.6666 | -104.6254
27.6702 | -104.6225
27.6739 | -104.622
27.6775 | -104.6229
27.6812 | -104.6226
27.6849 | -104.623
27.6885 | -104.6226
27.6922 | -104.6231
27.6958 | -104.6235
27.6995 | -104.6247
27.7031 | -104.6257
27.7068 | -104.6272
27.7105 | -104.6297
27.7141 | -104.6324
27.7178 | -104.6347
27.7214 | -104.6384
27.7251 | -104.6419
27.7287 | -104.6444
27.7324 | -104.6484
27.7361 | -104.6517
27.7397 | -104.656
27.7434 | -104.6593
27.747 | -104.6619
27.7507 | -104.6661
27.7544 | -104.6724
27.758 | -104.6778
27.7617 | -104.6835
27.7653 | -104.6884
27.769 | -104.693
27.7726 | -104.6985
27.7763 | -104.7035
27.78 | -104.7077
27.7836 | -104.712
27.7873 | -104.7161
27.7909 | -104.721
27.7946 | -104.7271
27.7983 | -104.7319
27.8019 | -104.7369
27.8056 | -104.743
27.8092 | -104.7486
27.8129 | -104.7543
27.8165 | -104.7598
27.8202 | -104.7644
27.8239 | -104.7691
27.8275 | -104.7758
27.8312 | -104.7819
27.8348 | -104.786
27.8385 | -104.7909
27.8422 | -104.7977
27.8458 | -104.8028
27.8495 | -104.8055
27.8531 | -104.8086
27.8568 | -104.8104
27.8604 | -104.8105
27.8641 | -104.8119
27.8678 | -104.816
27.8714 | -104.8211
27.8751 | -104.8235
27.8787 | -104.826
27.8824 | -104.8275
27.886 | -104.8286
27.8897 | -104.8296
27.8934 | -104.8297
27.897 | -104.831
27.9007 | -104.8319
27.9043 | -104.8317
27.908 | -104.8332
27.9117 | -104.8348
27.9153 | -104.8357
27.919 | -104.8343
27.9226 | -104.8316
27.9263 | -104.8285
27.9299 | -104.8255
27.9336 | -104.8217
27.9373 | -104.8187
27.9409 | -104.8154
27.9446 | -104.8108
27.9482 | -104.8059
27.9519 | -104.8026
27.9556 | -104.8009
27.9592 | -104.7972
27.9629 | -104.7942
27.9665 | -104.7922
27.9702 | -104.7865
27.9738 | -104.7824
27.9775 | -104.7772
27.9812 | -104.7723
27.9848 | -104.7672
27.9885 | -104.7622
27.9921 | -104.7559
27.9958 | -104.7496
27.9995 | -104.7444
28.0031 | -104.7383