## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 244 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 261
Hood | 02-92B/P | 180
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1005.5 | hPa |  | PRC PRO 01
Humidity | 56.4 | %/Hr
Internal temperature of calibration bench | 23.4 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1015 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.2 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.99 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3117 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 07/06/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 11/06/2018
Technical Manager | Anthony HUE | Approval |  | 13/06/2018
CEO | Robert GLEMOT | Approval |  | 13/06/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4809 | 57.6528
0.012209 | -36.0389 | 38.5102
0.018313 | -34.9725 | 28.3506
0.024417 | -34.5063 | 22.089
0.030521 | -34.3897 | 18.2868
0.036626 | -34.2437 | 15.2249
0.04273 | -34.1468 | 12.8656
0.048834 | -34.0986 | 11.3288
0.054938 | -34.0344 | 10.1737
0.061043 | -33.9752 | 9.0836
0.067147 | -33.9642 | 8.2606
0.073251 | -33.9641 | 7.5061
0.079355 | -33.9811 | 6.8561
0.08546 | -33.9944 | 6.273
0.091564 | -33.9891 | 5.8347
0.097668 | -33.9696 | 5.5776
0.10377 | -33.9469 | 5.32
0.10988 | -33.9294 | 5.0362
0.11598 | -33.9204 | 4.7406
0.12209 | -33.9226 | 4.4568
0.12819 | -33.933 | 4.1751
0.13429 | -33.9442 | 3.8897
0.1404 | -33.9478 | 3.6148
0.1465 | -33.95 | 3.3594
0.15261 | -33.9475 | 3.1394
0.15871 | -33.941 | 2.9736
0.16482 | -33.9324 | 2.8633
0.17092 | -33.9287 | 2.8166
0.17702 | -33.9256 | 2.739
0.18313 | -33.924 | 2.6567
0.18923 | -33.9235 | 2.5462
0.19534 | -33.9225 | 2.4192
0.20144 | -33.9218 | 2.2778
0.20754 | -33.9193 | 2.1569
0.21365 | -33.9186 | 2.0428
0.21975 | -33.9182 | 1.9454
0.22586 | -33.9193 | 1.8595
0.23196 | -33.9217 | 1.7888
0.23807 | -33.9256 | 1.7282
0.24417 | -33.9296 | 1.6713
0.25027 | -33.9334 | 1.6112
0.25638 | -33.9357 | 1.5501
0.26248 | -33.9376 | 1.4844
0.26859 | -33.9385 | 1.4197
0.27469 | -33.9389 | 1.349
0.2808 | -33.9399 | 1.2759
0.2869 | -33.9407 | 1.205
0.293 | -33.9415 | 1.1379
0.29911 | -33.9423 | 1.0775
0.30521 | -33.9431 | 1.0179
0.31132 | -33.9436 | 0.97901
0.31742 | -33.9445 | 0.93628
0.32353 | -33.9446 | 0.90128
0.32963 | -33.9452 | 0.87271
0.33573 | -33.9458 | 0.85311
0.34184 | -33.9468 | 0.83318
0.34794 | -33.9488 | 0.8122
0.35405 | -33.9506 | 0.7903
0.36015 | -33.9528 | 0.76734
0.36626 | -33.9546 | 0.7413
0.37236 | -33.9567 | 0.70648
0.37846 | -33.9589 | 0.66411
0.38457 | -33.96 | 0.60688
0.39067 | -33.9611 | 0.5604
0.39678 | -33.962 | 0.50832
0.40288 | -33.9627 | 0.45356
0.40899 | -33.9628 | 0.40027
0.41509 | -33.9625 | 0.34735
0.42119 | -33.9628 | 0.30238
0.4273 | -33.963 | 0.255
0.4334 | -33.9633 | 0.20783
0.43951 | -33.9638 | 0.16397
0.44561 | -33.9649 | 0.1248
0.45172 | -33.9665 | 0.095345
0.45782 | -33.9674 | 0.059555
0.46392 | -33.9687 | 0.025389
0.47003 | -33.9697 | -0.0047659
0.47613 | -33.9706 | -0.036824
0.48224 | -33.9712 | -0.069484
0.48834 | -33.9703 | -0.10403
0.49445 | -33.9702 | -0.13831
0.50055 | -33.9697 | -0.17013
0.50665 | -33.969 | -0.20462
0.51276 | -33.9682 | -0.2394
0.51886 | -33.9674 | -0.27426
0.52497 | -33.9663 | -0.30121
0.53107 | -33.9657 | -0.33236
0.53717 | -33.9652 | -0.3622
0.54328 | -33.9644 | -0.38724
0.54938 | -33.9643 | -0.41155
0.55549 | -33.9643 | -0.43416
0.56159 | -33.9652 | -0.44664
0.5677 | -33.9654 | -0.46299
0.5738 | -33.9664 | -0.47979
0.5799 | -33.9671 | -0.49535
0.58601 | -33.9679 | -0.50579
0.59211 | -33.9691 | -0.51131
0.59822 | -33.9702 | -0.52047
0.60432 | -33.9714 | -0.53197
0.61043 | -33.9723 | -0.53923
0.61653 | -33.9731 | -0.5488
0.62263 | -33.9743 | -0.55741
0.62874 | -33.9745 | -0.5676
0.63484 | -33.975 | -0.57646
0.64095 | -33.9751 | -0.58623
0.64705 | -33.9752 | -0.59578
0.65316 | -33.975 | -0.60786
0.65926 | -33.9749 | -0.62301
0.66536 | -33.9737 | -0.64045
0.67147 | -33.9731 | -0.65376
0.67757 | -33.9726 | -0.66811
0.68368 | -33.9718 | -0.68687
0.68978 | -33.9709 | -0.70399
0.69589 | -33.97 | -0.72297
0.70199 | -33.9693 | -0.7456
0.70809 | -33.9684 | -0.76455
0.7142 | -33.9678 | -0.78488
0.7203 | -33.9668 | -0.80424
0.72641 | -33.9663 | -0.82676
0.73251 | -33.9658 | -0.84682
0.73862 | -33.9663 | -0.86864
0.74472 | -33.9661 | -0.88889
0.75082 | -33.966 | -0.91111
0.75693 | -33.9665 | -0.93302
0.76303 | -33.9671 | -0.955
0.76914 | -33.969 | -0.97537
0.77524 | -33.97 | -0.99504
0.78135 | -33.9713 | -1.015
0.78745 | -33.9728 | -1.035
0.79355 | -33.974 | -1.0548
0.79966 | -33.9758 | -1.0742
0.80576 | -33.9779 | -1.0939
0.81187 | -33.9795 | -1.1154
0.81797 | -33.9814 | -1.1366
0.82408 | -33.9831 | -1.1569
0.83018 | -33.9847 | -1.1777
0.83628 | -33.9865 | -1.1963
0.84239 | -33.9875 | -1.2169
0.84849 | -33.9889 | -1.2385
0.8546 | -33.9903 | -1.2584
0.8607 | -33.9914 | -1.2795
0.86681 | -33.9924 | -1.2989
0.87291 | -33.9933 | -1.3181
0.87901 | -33.9936 | -1.3367
0.88512 | -33.9943 | -1.3552
0.89122 | -33.995 | -1.3737
0.89733 | -33.9953 | -1.3897
0.90343 | -33.9956 | -1.4077
0.90953 | -33.9957 | -1.4227
0.91564 | -33.9957 | -1.439
0.92174 | -33.9958 | -1.4544
0.92785 | -33.9958 | -1.4696
0.93395 | -33.9959 | -1.4838
0.94006 | -33.996 | -1.4973
0.94616 | -33.9956 | -1.5094
0.95226 | -33.9957 | -1.5226
0.95837 | -33.9956 | -1.5369
0.96447 | -33.9957 | -1.5501
0.97058 | -33.9955 | -1.5637
0.97668 | -33.9952 | -1.5772
0.98279 | -33.9949 | -1.5897
0.98889 | -33.9946 | -1.6006
0.99499 | -33.9945 | -1.6137
1.0011 | -33.9943 | -1.6245
1.0072 | -33.9939 | -1.6361
1.0133 | -33.9935 | -1.6467
1.0194 | -33.993 | -1.6592
1.0255 | -33.9926 | -1.6726
1.0316 | -33.992 | -1.684
1.0377 | -33.9916 | -1.6971
1.0438 | -33.9909 | -1.7092
1.0499 | -33.9901 | -1.7235
1.056 | -33.9893 | -1.7372
1.0621 | -33.9886 | -1.7508
1.0682 | -33.9881 | -1.7626
1.0743 | -33.9873 | -1.7771
1.0805 | -33.9868 | -1.7908
1.0866 | -33.9858 | -1.8047
1.0927 | -33.9849 | -1.8183
1.0988 | -33.9842 | -1.8323
1.1049 | -33.9834 | -1.8456
1.111 | -33.9825 | -1.86
1.1171 | -33.9818 | -1.8736
1.1232 | -33.9813 | -1.888
1.1293 | -33.9808 | -1.9013
1.1354 | -33.9805 | -1.9155
1.1415 | -33.98 | -1.9303
1.1476 | -33.9794 | -1.9449
1.1537 | -33.9788 | -1.9588
1.1598 | -33.9789 | -1.9731
1.1659 | -33.9786 | -1.9857
1.172 | -33.9783 | -1.9989
1.1781 | -33.9781 | -2.0121
1.1842 | -33.9779 | -2.0256
1.1903 | -33.9783 | -2.0367
1.1964 | -33.9782 | -2.0496
1.2025 | -33.9783 | -2.0633
1.2086 | -33.9786 | -2.0755
1.2147 | -33.9787 | -2.0876
1.2209 | -33.9791 | -2.0987
1.227 | -33.9798 | -2.1098
1.2331 | -33.9803 | -2.1219
1.2392 | -33.9807 | -2.1338
1.2453 | -33.9812 | -2.1449
1.2514 | -33.9816 | -2.1548
1.2575 | -33.9823 | -2.1662
1.2636 | -33.9832 | -2.1758
1.2697 | -33.9839 | -2.1854
1.2758 | -33.9846 | -2.1968
1.2819 | -33.9854 | -2.2073
1.288 | -33.9861 | -2.2184
1.2941 | -33.987 | -2.2274
1.3002 | -33.9878 | -2.2361
1.3063 | -33.9885 | -2.2456
1.3124 | -33.9892 | -2.2551
1.3185 | -33.99 | -2.2646
1.3246 | -33.9908 | -2.2742
1.3307 | -33.9918 | -2.2834
1.3368 | -33.9926 | -2.2928
1.3429 | -33.9934 | -2.3023
1.349 | -33.994 | -2.312
1.3551 | -33.9947 | -2.3218
1.3613 | -33.9954 | -2.3305
1.3674 | -33.9961 | -2.341
1.3735 | -33.9966 | -2.351
1.3796 | -33.9971 | -2.3603
1.3857 | -33.9975 | -2.3697
1.3918 | -33.9981 | -2.3801
1.3979 | -33.9987 | -2.3897
1.404 | -33.9992 | -2.4013
1.4101 | -33.9996 | -2.4117
1.4162 | -34 | -2.4225
1.4223 | -34.0005 | -2.4341
1.4284 | -34.001 | -2.4444
1.4345 | -34.0013 | -2.4557
1.4406 | -34.0015 | -2.4691
1.4467 | -34.002 | -2.4808
1.4528 | -34.0024 | -2.4923
1.4589 | -34.0029 | -2.5042
1.465 | -34.0031 | -2.5149
1.4711 | -34.0036 | -2.5264
1.4772 | -34.0041 | -2.5399
1.4833 | -34.0045 | -2.553
1.4894 | -34.0049 | -2.5653
1.4955 | -34.0055 | -2.5775
1.5016 | -34.006 | -2.5902
1.5078 | -34.0062 | -2.6041
1.5139 | -34.0066 | -2.6171
1.52 | -34.007 | -2.6298
1.5261 | -34.0072 | -2.6424
1.5322 | -34.0075 | -2.6563
1.5383 | -34.0078 | -2.669
1.5444 | -34.0083 | -2.6818
1.5505 | -34.0087 | -2.6954
1.5566 | -34.0092 | -2.7087
1.5627 | -34.0095 | -2.7225
1.5688 | -34.0101 | -2.7363
1.5749 | -34.0104 | -2.7489
1.581 | -34.0106 | -2.762
1.5871 | -34.0109 | -2.7757
1.5932 | -34.0112 | -2.7885
1.5993 | -34.0115 | -2.8027
1.6054 | -34.0118 | -2.8158
1.6115 | -34.0122 | -2.83
1.6176 | -34.0124 | -2.8424
1.6237 | -34.0125 | -2.8558
1.6298 | -34.0128 | -2.8682
1.6359 | -34.0131 | -2.8809
1.642 | -34.0135 | -2.8929
1.6482 | -34.0136 | -2.9057
1.6543 | -34.0139 | -2.9193
1.6604 | -34.0142 | -2.9313
1.6665 | -34.0143 | -2.9432
1.6726 | -34.0145 | -2.9555
1.6787 | -34.0148 | -2.968
1.6848 | -34.0146 | -2.9799
1.6909 | -34.0148 | -2.9924
1.697 | -34.015 | -3.0045
1.7031 | -34.0151 | -3.0166
1.7092 | -34.015 | -3.0282
1.7153 | -34.0151 | -3.0399
1.7214 | -34.0149 | -3.0515
1.7275 | -34.0149 | -3.0634
1.7336 | -34.0148 | -3.0759
1.7397 | -34.0149 | -3.088
1.7458 | -34.0149 | -3.0997
1.7519 | -34.0149 | -3.1114
1.758 | -34.0149 | -3.123
1.7641 | -34.015 | -3.1344
1.7702 | -34.0149 | -3.1468
1.7763 | -34.0149 | -3.1578
1.7824 | -34.0149 | -3.1698
1.7885 | -34.0148 | -3.1818
1.7947 | -34.0149 | -3.1937
1.8008 | -34.015 | -3.2046
1.8069 | -34.015 | -3.2158
1.813 | -34.015 | -3.2277
1.8191 | -34.015 | -3.24
1.8252 | -34.015 | -3.2525
1.8313 | -34.0151 | -3.2646
1.8374 | -34.0153 | -3.2763
1.8435 | -34.0153 | -3.2881
1.8496 | -34.0153 | -3.3009
1.8557 | -34.0155 | -3.3135
1.8618 | -34.0157 | -3.3259
1.8679 | -34.0158 | -3.3378
1.874 | -34.0159 | -3.3497
1.8801 | -34.0161 | -3.3621
1.8862 | -34.0163 | -3.3741
1.8923 | -34.0164 | -3.3862
1.8984 | -34.0165 | -3.3982
1.9045 | -34.0166 | -3.4098
1.9106 | -34.0167 | -3.4218
1.9167 | -34.0167 | -3.4338
1.9228 | -34.0169 | -3.4457
1.9289 | -34.017 | -3.4575
1.9351 | -34.0172 | -3.4685
1.9412 | -34.0174 | -3.4798
1.9473 | -34.0177 | -3.492
1.9534 | -34.0179 | -3.5035
1.9595 | -34.0181 | -3.5153
1.9656 | -34.0183 | -3.5277
1.9717 | -34.0186 | -3.5395
1.9778 | -34.0189 | -3.5512
1.9839 | -34.0192 | -3.5627
1.99 | -34.0194 | -3.5739
1.9961 | -34.0195 | -3.5859
2.0022 | -34.0197 | -3.5974
2.0083 | -34.0198 | -3.6085
2.0144 | -34.02 | -3.6199
2.0205 | -34.0203 | -3.6318
2.0266 | -34.0207 | -3.6434
2.0327 | -34.0209 | -3.6551
2.0388 | -34.0213 | -3.6668
2.0449 | -34.0215 | -3.6782
2.051 | -34.0218 | -3.6898
2.0571 | -34.0221 | -3.7012
2.0632 | -34.0223 | -3.7127
2.0693 | -34.0225 | -3.7248
2.0754 | -34.0227 | -3.7361
2.0816 | -34.023 | -3.7482
2.0877 | -34.0232 | -3.7597
2.0938 | -34.0236 | -3.772
2.0999 | -34.0237 | -3.7835
2.106 | -34.0239 | -3.795
2.1121 | -34.0241 | -3.8069
2.1182 | -34.0244 | -3.8177
2.1243 | -34.0246 | -3.8285
2.1304 | -34.0248 | -3.8397
2.1365 | -34.0251 | -3.8514
2.1426 | -34.0253 | -3.8632
2.1487 | -34.0256 | -3.8745
2.1548 | -34.026 | -3.8865
2.1609 | -34.0264 | -3.898
2.167 | -34.0268 | -3.9093
2.1731 | -34.0271 | -3.9213
2.1792 | -34.0273 | -3.9323
2.1853 | -34.0276 | -3.9444
2.1914 | -34.0279 | -3.9558
2.1975 | -34.0283 | -3.9674
2.2036 | -34.0285 | -3.9795
2.2097 | -34.0287 | -3.9909
2.2158 | -34.0291 | -4.0016
2.222 | -34.0293 | -4.0121
2.2281 | -34.0296 | -4.0235
2.2342 | -34.0299 | -4.0349
2.2403 | -34.0301 | -4.0455
2.2464 | -34.0304 | -4.0562
2.2525 | -34.0306 | -4.0677
2.2586 | -34.0309 | -4.0785
2.2647 | -34.0312 | -4.0891
2.2708 | -34.0316 | -4.1003
2.2769 | -34.0319 | -4.1112
2.283 | -34.0321 | -4.1223
2.2891 | -34.0323 | -4.1332
2.2952 | -34.0325 | -4.1437
2.3013 | -34.0328 | -4.1545
2.3074 | -34.033 | -4.1649
2.3135 | -34.0332 | -4.1763
2.3196 | -34.0335 | -4.1866
2.3257 | -34.0337 | -4.1974
2.3318 | -34.0339 | -4.2085
2.3379 | -34.0341 | -4.2191
2.344 | -34.0342 | -4.2305
2.3501 | -34.0344 | -4.2411
2.3562 | -34.0346 | -4.2516
2.3623 | -34.0348 | -4.262
2.3685 | -34.0349 | -4.2725
2.3746 | -34.035 | -4.2836
2.3807 | -34.0353 | -4.2943
2.3868 | -34.0355 | -4.3051
2.3929 | -34.0357 | -4.3158
2.399 | -34.0358 | -4.3262
2.4051 | -34.0361 | -4.3369
2.4112 | -34.0363 | -4.3476
2.4173 | -34.0365 | -4.3582
2.4234 | -34.0365 | -4.3691
2.4295 | -34.0367 | -4.3798
2.4356 | -34.0369 | -4.3908
2.4417 | -34.037 | -4.402
2.4478 | -34.0372 | -4.4127
2.4539 | -34.0374 | -4.4231
2.46 | -34.0375 | -4.4344
2.4661 | -34.0376 | -4.4456
2.4722 | -34.0378 | -4.4559
2.4783 | -34.0379 | -4.4664
2.4844 | -34.038 | -4.4773
2.4905 | -34.0381 | -4.4883
2.4966 | -34.0383 | -4.4991
2.5027 | -34.0384 | -4.5095
2.5089 | -34.0385 | -4.5204
2.515 | -34.0385 | -4.5311
2.5211 | -34.0386 | -4.5418
2.5272 | -34.0387 | -4.5526
2.5333 | -34.0388 | -4.5636
2.5394 | -34.039 | -4.5745
2.5455 | -34.0391 | -4.5859
2.5516 | -34.0393 | -4.5962
2.5577 | -34.0394 | -4.6077
2.5638 | -34.0396 | -4.6193
2.5699 | -34.0396 | -4.6306
2.576 | -34.0397 | -4.6421
2.5821 | -34.0398 | -4.6534
2.5882 | -34.0399 | -4.6644
2.5943 | -34.04 | -4.6753
2.6004 | -34.0402 | -4.6862
2.6065 | -34.0403 | -4.6965
2.6126 | -34.0404 | -4.707
2.6187 | -34.0405 | -4.7173
2.6248 | -34.0407 | -4.728
2.6309 | -34.0409 | -4.7394
2.637 | -34.0411 | -4.7509
2.6431 | -34.0412 | -4.7612
2.6492 | -34.0413 | -4.7725
2.6554 | -34.0415 | -4.7833
2.6615 | -34.0416 | -4.7941
2.6676 | -34.0418 | -4.8055
2.6737 | -34.0421 | -4.8166
2.6798 | -34.0423 | -4.8277
2.6859 | -34.0425 | -4.8387
2.692 | -34.0426 | -4.8495
2.6981 | -34.0428 | -4.8606
2.7042 | -34.043 | -4.8717
2.7103 | -34.0431 | -4.8824
2.7164 | -34.0434 | -4.8933
2.7225 | -34.0436 | -4.9045
2.7286 | -34.0438 | -4.9155
2.7347 | -34.044 | -4.926
2.7408 | -34.0443 | -4.9361
2.7469 | -34.0444 | -4.9474
2.753 | -34.0447 | -4.9579
2.7591 | -34.045 | -4.9692
2.7652 | -34.0452 | -4.9805
2.7713 | -34.0453 | -4.9912
2.7774 | -34.0456 | -5.0021
2.7835 | -34.0457 | -5.013
2.7896 | -34.0459 | -5.0234
2.7958 | -34.0462 | -5.0341
2.8019 | -34.0465 | -5.0444
2.808 | -34.0467 | -5.0552
2.8141 | -34.047 | -5.0663
2.8202 | -34.0473 | -5.0771
2.8263 | -34.0476 | -5.0879
2.8324 | -34.0479 | -5.0992
2.8385 | -34.0483 | -5.1101
2.8446 | -34.0486 | -5.1208
2.8507 | -34.049 | -5.1317
2.8568 | -34.0493 | -5.1422
2.8629 | -34.0495 | -5.1527
2.869 | -34.0498 | -5.1638
2.8751 | -34.0501 | -5.1745
2.8812 | -34.0505 | -5.1857
2.8873 | -34.0508 | -5.1966
2.8934 | -34.0511 | -5.2075
2.8995 | -34.0515 | -5.2185
2.9056 | -34.0517 | -5.2296
2.9117 | -34.052 | -5.2406
2.9178 | -34.0524 | -5.2514
2.9239 | -34.0528 | -5.2629
2.93 | -34.0532 | -5.2744
2.9361 | -34.0535 | -5.2852
2.9423 | -34.0538 | -5.296
2.9484 | -34.0542 | -5.3071
2.9545 | -34.0546 | -5.3181
2.9606 | -34.055 | -5.3294
2.9667 | -34.0553 | -5.3407
2.9728 | -34.0556 | -5.352
2.9789 | -34.0561 | -5.3631
2.985 | -34.0565 | -5.3744
2.9911 | -34.0569 | -5.3857
2.9972 | -34.0573 | -5.3968
3.0033 | -34.0577 | -5.4083
3.0094 | -34.0581 | -5.4194
3.0155 | -34.0585 | -5.4304
3.0216 | -34.0589 | -5.4417
3.0277 | -34.0592 | -5.4534
3.0338 | -34.0596 | -5.4647
3.0399 | -34.06 | -5.4759
3.046 | -34.0603 | -5.4875
3.0521 | -34.0607 | -5.4987
3.0582 | -34.061 | -5.5109
3.0643 | -34.0613 | -5.5224
3.0704 | -34.0617 | -5.5342
3.0765 | -34.062 | -5.5453
3.0827 | -34.0623 | -5.5565
3.0888 | -34.0626 | -5.5678
3.0949 | -34.063 | -5.5791
3.101 | -34.0634 | -5.5907
3.1071 | -34.0637 | -5.6025
3.1132 | -34.0641 | -5.6143
3.1193 | -34.0644 | -5.6257
3.1254 | -34.0649 | -5.637
3.1315 | -34.0652 | -5.6489
3.1376 | -34.0655 | -5.6603
3.1437 | -34.0659 | -5.6719
3.1498 | -34.0662 | -5.684
3.1559 | -34.0665 | -5.6954
3.162 | -34.0668 | -5.7075
3.1681 | -34.0671 | -5.7195
3.1742 | -34.0674 | -5.7311
3.1803 | -34.0678 | -5.7428
3.1864 | -34.0681 | -5.7541
3.1925 | -34.0684 | -5.7657
3.1986 | -34.0687 | -5.7777
3.2047 | -34.069 | -5.7897
3.2108 | -34.0693 | -5.801
3.2169 | -34.0696 | -5.8127
3.223 | -34.0699 | -5.8245
3.2292 | -34.0701 | -5.8364
3.2353 | -34.0703 | -5.8473
3.2414 | -34.0706 | -5.8589
3.2475 | -34.0708 | -5.8703
3.2536 | -34.0711 | -5.8818
3.2597 | -34.0714 | -5.8928
3.2658 | -34.0716 | -5.9038
3.2719 | -34.0718 | -5.9149
3.278 | -34.0721 | -5.9266
3.2841 | -34.0723 | -5.9378
3.2902 | -34.0726 | -5.9492
3.2963 | -34.0728 | -5.9608
3.3024 | -34.0729 | -5.972
3.3085 | -34.0732 | -5.9836
3.3146 | -34.0734 | -5.9948
3.3207 | -34.0737 | -6.0063
3.3268 | -34.0739 | -6.0179
3.3329 | -34.0741 | -6.0288
3.339 | -34.0742 | -6.0397
3.3451 | -34.0744 | -6.051
3.3512 | -34.0746 | -6.0618
3.3573 | -34.0748 | -6.0731
3.3634 | -34.075 | -6.0843
3.3696 | -34.0752 | -6.0958
3.3757 | -34.0753 | -6.1067
3.3818 | -34.0756 | -6.1176
3.3879 | -34.0757 | -6.1288
3.394 | -34.0759 | -6.1402
3.4001 | -34.0761 | -6.1512
3.4062 | -34.0762 | -6.1622
3.4123 | -34.0763 | -6.1729
3.4184 | -34.0765 | -6.1842
3.4245 | -34.0767 | -6.1951
3.4306 | -34.0768 | -6.2059
3.4367 | -34.0769 | -6.2165
3.4428 | -34.0771 | -6.227
3.4489 | -34.0772 | -6.238
3.455 | -34.0773 | -6.249
3.4611 | -34.0775 | -6.2603
3.4672 | -34.0777 | -6.2714
3.4733 | -34.0778 | -6.2826
3.4794 | -34.078 | -6.2934
3.4855 | -34.0781 | -6.3039
3.4916 | -34.0783 | -6.3144
3.4977 | -34.0784 | -6.3257
3.5038 | -34.0786 | -6.3367
3.5099 | -34.0787 | -6.3471
3.5161 | -34.0788 | -6.3576
3.5222 | -34.0789 | -6.368
3.5283 | -34.079 | -6.379
3.5344 | -34.0791 | -6.3898
3.5405 | -34.0793 | -6.4004
3.5466 | -34.0794 | -6.4114
3.5527 | -34.0795 | -6.4223
3.5588 | -34.0796 | -6.433
3.5649 | -34.0797 | -6.4438
3.571 | -34.0798 | -6.4545
3.5771 | -34.0799 | -6.4657
3.5832 | -34.08 | -6.4756
3.5893 | -34.0802 | -6.4863
3.5954 | -34.0804 | -6.4972
3.6015 | -34.0805 | -6.5081
3.6076 | -34.0807 | -6.5187
3.6137 | -34.0808 | -6.5289
3.6198 | -34.0811 | -6.5391
3.6259 | -34.0812 | -6.5501
3.632 | -34.0814 | -6.5607
3.6381 | -34.0815 | -6.5716
3.6442 | -34.0817 | -6.5822
3.6503 | -34.0818 | -6.5924
3.6565 | -34.082 | -6.6032
3.6626 | -34.0821 | -6.6138
3.6687 | -34.0824 | -6.6246
3.6748 | -34.0826 | -6.635
3.6809 | -34.0828 | -6.6456
3.687 | -34.083 | -6.6566
3.6931 | -34.0832 | -6.6668
3.6992 | -34.0834 | -6.6775
3.7053 | -34.0836 | -6.6879
3.7114 | -34.0837 | -6.6987
3.7175 | -34.0839 | -6.7094
3.7236 | -34.0842 | -6.7194
3.7297 | -34.0844 | -6.7302
3.7358 | -34.0846 | -6.7411
3.7419 | -34.0847 | -6.752
3.748 | -34.085 | -6.7625
3.7541 | -34.0851 | -6.7731
3.7602 | -34.0854 | -6.7835
3.7663 | -34.0856 | -6.794
3.7724 | -34.0858 | -6.8046
3.7785 | -34.0861 | -6.8149
3.7846 | -34.0863 | -6.8254
3.7907 | -34.0866 | -6.8362
3.7969 | -34.0868 | -6.8469
3.803 | -34.0871 | -6.8571
3.8091 | -34.0873 | -6.8676
3.8152 | -34.0875 | -6.878
3.8213 | -34.0878 | -6.8889
3.8274 | -34.088 | -6.8991
3.8335 | -34.0883 | -6.9098
3.8396 | -34.0886 | -6.9204
3.8457 | -34.0889 | -6.931
3.8518 | -34.0892 | -6.9414
3.8579 | -34.0896 | -6.9519
3.864 | -34.0898 | -6.9626
3.8701 | -34.0901 | -6.9735
3.8762 | -34.0904 | -6.9839
3.8823 | -34.0906 | -6.9947
3.8884 | -34.0909 | -7.0053
3.8945 | -34.0913 | -7.0161
3.9006 | -34.0916 | -7.027
3.9067 | -34.0919 | -7.038
3.9128 | -34.0921 | -7.0487
3.9189 | -34.0924 | -7.059
3.925 | -34.0927 | -7.0698
3.9311 | -34.0929 | -7.08
3.9372 | -34.0933 | -7.0907
3.9434 | -34.0935 | -7.1014
3.9495 | -34.0939 | -7.1121
3.9556 | -34.0942 | -7.1225
3.9617 | -34.0945 | -7.1328
3.9678 | -34.0948 | -7.143
3.9739 | -34.0951 | -7.1536
3.98 | -34.0954 | -7.1644
3.9861 | -34.0957 | -7.1749
3.9922 | -34.0959 | -7.1856
3.9983 | -34.0962 | -7.1957
4.0044 | -34.0965 | -7.2059
4.0105 | -34.0968 | -7.2165
4.0166 | -34.0971 | -7.2271
4.0227 | -34.0974 | -7.2377
4.0288 | -34.0976 | -7.2487
4.0349 | -34.0979 | -7.2591
4.041 | -34.0981 | -7.2696
4.0471 | -34.0983 | -7.2804
4.0532 | -34.0987 | -7.2906
4.0593 | -34.099 | -7.3014
4.0654 | -34.0993 | -7.3117
4.0715 | -34.0995 | -7.3227
4.0776 | -34.0999 | -7.3332
4.0838 | -34.1002 | -7.3442
4.0899 | -34.1004 | -7.3545
4.096 | -34.1007 | -7.365
4.1021 | -34.101 | -7.3754
4.1082 | -34.1012 | -7.3858
4.1143 | -34.1015 | -7.396
4.1204 | -34.1018 | -7.4063
4.1265 | -34.1021 | -7.4166
4.1326 | -34.1023 | -7.4268
4.1387 | -34.1027 | -7.4372
4.1448 | -34.1029 | -7.4478
4.1509 | -34.1031 | -7.4583
4.157 | -34.1034 | -7.4692
4.1631 | -34.1037 | -7.4799
4.1692 | -34.1039 | -7.4902
4.1753 | -34.1042 | -7.5007
4.1814 | -34.1044 | -7.5109
4.1875 | -34.1047 | -7.5216
4.1936 | -34.105 | -7.5322
4.1997 | -34.1052 | -7.5424
4.2058 | -34.1054 | -7.5531
4.2119 | -34.1057 | -7.5635
4.218 | -34.1059 | -7.5738
4.2241 | -34.1061 | -7.5841
4.2303 | -34.1064 | -7.5944
4.2364 | -34.1066 | -7.6048
4.2425 | -34.1068 | -7.6152
4.2486 | -34.107 | -7.6259
4.2547 | -34.1072 | -7.6361
4.2608 | -34.1074 | -7.6467
4.2669 | -34.1076 | -7.6568
4.273 | -34.1078 | -7.6675
4.2791 | -34.1081 | -7.6779
4.2852 | -34.1083 | -7.6885
4.2913 | -34.1084 | -7.6986
4.2974 | -34.1087 | -7.7093
4.3035 | -34.1089 | -7.72
4.3096 | -34.1091 | -7.7306
4.3157 | -34.1094 | -7.7408
4.3218 | -34.1096 | -7.7512
4.3279 | -34.1098 | -7.7617
4.334 | -34.11 | -7.772
4.3401 | -34.1102 | -7.782
4.3462 | -34.1104 | -7.7925
4.3523 | -34.1107 | -7.8024
4.3584 | -34.1108 | -7.8125
4.3645 | -34.111 | -7.8227
4.3707 | -34.1112 | -7.8328
4.3768 | -34.1114 | -7.8431
4.3829 | -34.1117 | -7.8538
4.389 | -34.1119 | -7.8642
4.3951 | -34.1121 | -7.8744
4.4012 | -34.1122 | -7.885
4.4073 | -34.1124 | -7.8952
4.4134 | -34.1126 | -7.9055
4.4195 | -34.1129 | -7.9157
4.4256 | -34.1131 | -7.9258
4.4317 | -34.1133 | -7.9361
4.4378 | -34.1135 | -7.9461
4.4439 | -34.1137 | -7.9562
4.45 | -34.1138 | -7.9665
4.4561 | -34.114 | -7.9768
4.4622 | -34.1142 | -7.9872
4.4683 | -34.1144 | -7.9974
4.4744 | -34.1146 | -8.008
4.4805 | -34.1148 | -8.0179
4.4866 | -34.115 | -8.0281
4.4927 | -34.1152 | -8.0386
4.4988 | -34.1154 | -8.0489
4.5049 | -34.1156 | -8.0588
4.511 | -34.1158 | -8.0687
4.5172 | -34.116 | -8.0789
4.5233 | -34.1162 | -8.0896
4.5294 | -34.1164 | -8.0996
4.5355 | -34.1166 | -8.1099
4.5416 | -34.1168 | -8.1201
4.5477 | -34.117 | -8.1304
4.5538 | -34.1171 | -8.1402
4.5599 | -34.1174 | -8.1504
4.566 | -34.1175 | -8.1606
4.5721 | -34.1177 | -8.1709
4.5782 | -34.118 | -8.1813
4.5843 | -34.1182 | -8.1916
4.5904 | -34.1184 | -8.2019
4.5965 | -34.1187 | -8.212
4.6026 | -34.1189 | -8.2226
4.6087 | -34.119 | -8.2328
4.6148 | -34.1192 | -8.2434
4.6209 | -34.1194 | -8.2535
4.627 | -34.1197 | -8.264
4.6331 | -34.1198 | -8.2741
4.6392 | -34.12 | -8.2844
4.6453 | -34.1202 | -8.2946
4.6514 | -34.1205 | -8.305
4.6576 | -34.1206 | -8.3156
4.6637 | -34.1208 | -8.3255
4.6698 | -34.121 | -8.3355
4.6759 | -34.1212 | -8.3461
4.682 | -34.1214 | -8.3563
4.6881 | -34.1215 | -8.3666
4.6942 | -34.1217 | -8.3768
4.7003 | -34.1219 | -8.3871
4.7064 | -34.1221 | -8.3971
4.7125 | -34.1224 | -8.4076
4.7186 | -34.1225 | -8.4176
4.7247 | -34.1228 | -8.4279
4.7308 | -34.123 | -8.4381
4.7369 | -34.1232 | -8.4485
4.743 | -34.1234 | -8.4591
4.7491 | -34.1236 | -8.4697
4.7552 | -34.1238 | -8.4801
4.7613 | -34.124 | -8.4905
4.7674 | -34.1242 | -8.5006
4.7735 | -34.1244 | -8.5113
4.7796 | -34.1246 | -8.5218
4.7857 | -34.1248 | -8.5323
4.7918 | -34.125 | -8.5424
4.7979 | -34.1253 | -8.5526
4.8041 | -34.1255 | -8.563
4.8102 | -34.1257 | -8.573
4.8163 | -34.1259 | -8.5833
4.8224 | -34.1261 | -8.5937
4.8285 | -34.1264 | -8.6037
4.8346 | -34.1266 | -8.6138
4.8407 | -34.1268 | -8.6241
4.8468 | -34.127 | -8.6347
4.8529 | -34.1273 | -8.6453
4.859 | -34.1275 | -8.6559
4.8651 | -34.1277 | -8.666
4.8712 | -34.1279 | -8.6762
4.8773 | -34.1281 | -8.6866
4.8834 | -34.1283 | -8.6969
4.8895 | -34.1286 | -8.707
4.8956 | -34.1287 | -8.7173
4.9017 | -34.1289 | -8.7275
4.9078 | -34.1291 | -8.7381
4.9139 | -34.1294 | -8.748
4.92 | -34.1296 | -8.7581
4.9261 | -34.1299 | -8.768
4.9322 | -34.1301 | -8.7784
4.9383 | -34.1304 | -8.7887
4.9445 | -34.1306 | -8.799
4.9506 | -34.1308 | -8.8092
4.9567 | -34.131 | -8.8192
4.9628 | -34.1312 | -8.8295
4.9689 | -34.1315 | -8.8401
4.975 | -34.1317 | -8.8507
4.9811 | -34.1319 | -8.8608
4.9872 | -34.1322 | -8.8709
4.9933 | -34.1325 | -8.8815
4.9994 | -34.1327 | -8.8922
5.0055 | -34.133 | -8.9025
5.0116 | -34.1332 | -8.9133
5.0177 | -34.1335 | -8.9238
5.0238 | -34.1338 | -8.9344
5.0299 | -34.1341 | -8.9451
5.036 | -34.1343 | -8.9554
5.0421 | -34.1345 | -8.9656
5.0482 | -34.1348 | -8.9761
5.0543 | -34.135 | -8.9861
5.0604 | -34.1353 | -8.9965
5.0665 | -34.1356 | -9.0065
5.0726 | -34.1358 | -9.0171
5.0787 | -34.1361 | -9.0277
5.0848 | -34.1364 | -9.0381
5.091 | -34.1366 | -9.0482
5.0971 | -34.1369 | -9.0583
5.1032 | -34.1372 | -9.069
5.1093 | -34.1374 | -9.0794
5.1154 | -34.1376 | -9.0895
5.1215 | -34.1379 | -9.0996
5.1276 | -34.1381 | -9.1098
5.1337 | -34.1385 | -9.1199
5.1398 | -34.1387 | -9.13
5.1459 | -34.139 | -9.1406
5.152 | -34.1392 | -9.1511
5.1581 | -34.1395 | -9.1616
5.1642 | -34.1397 | -9.172
5.1703 | -34.14 | -9.1823
5.1764 | -34.1403 | -9.1925
5.1825 | -34.1406 | -9.2026
5.1886 | -34.1408 | -9.2127
5.1947 | -34.1411 | -9.2231
5.2008 | -34.1414 | -9.2335
5.2069 | -34.1416 | -9.2441
5.213 | -34.1419 | -9.2543
5.2191 | -34.1422 | -9.2646
5.2252 | -34.1425 | -9.275
5.2314 | -34.1427 | -9.2856
5.2375 | -34.143 | -9.2957
5.2436 | -34.1432 | -9.3061
5.2497 | -34.1435 | -9.3162
5.2558 | -34.1438 | -9.3269
5.2619 | -34.1441 | -9.3375
5.268 | -34.1444 | -9.348
5.2741 | -34.1447 | -9.3584
5.2802 | -34.1449 | -9.3692
5.2863 | -34.1453 | -9.3794
5.2924 | -34.1455 | -9.39
5.2985 | -34.1458 | -9.4005
5.3046 | -34.1461 | -9.4109
5.3107 | -34.1464 | -9.4215
5.3168 | -34.1467 | -9.4319
5.3229 | -34.147 | -9.4423
5.329 | -34.1472 | -9.4527
5.3351 | -34.1475 | -9.4628
5.3412 | -34.1478 | -9.4732
5.3473 | -34.1481 | -9.4836
5.3534 | -34.1484 | -9.4942
5.3595 | -34.1487 | -9.5047
5.3656 | -34.149 | -9.515
5.3717 | -34.1492 | -9.5258
5.3779 | -34.1495 | -9.5365
5.384 | -34.1498 | -9.5469
5.3901 | -34.1501 | -9.5576
5.3962 | -34.1504 | -9.5679
5.4023 | -34.1507 | -9.5785
5.4084 | -34.151 | -9.5889
5.4145 | -34.1514 | -9.599
5.4206 | -34.1516 | -9.6095
5.4267 | -34.1519 | -9.6203
5.4328 | -34.1522 | -9.6308
5.4389 | -34.1525 | -9.641
5.445 | -34.1528 | -9.6518
5.4511 | -34.1531 | -9.662
5.4572 | -34.1534 | -9.6725
5.4633 | -34.1537 | -9.6829
5.4694 | -34.1539 | -9.6934
5.4755 | -34.1542 | -9.7036
5.4816 | -34.1545 | -9.7141
5.4877 | -34.1549 | -9.7244
5.4938 | -34.1551 | -9.7346
5.4999 | -34.1555 | -9.7449
5.506 | -34.1558 | -9.7553
5.5121 | -34.1561 | -9.7656
5.5183 | -34.1564 | -9.7759
5.5244 | -34.1566 | -9.7866
5.5305 | -34.157 | -9.7969
5.5366 | -34.1572 | -9.8071
5.5427 | -34.1575 | -9.8177
5.5488 | -34.1578 | -9.8281
5.5549 | -34.1581 | -9.8386
5.561 | -34.1584 | -9.8493
5.5671 | -34.1587 | -9.8596
5.5732 | -34.159 | -9.8703
5.5793 | -34.1593 | -9.881
5.5854 | -34.1596 | -9.8914
5.5915 | -34.1599 | -9.9021
5.5976 | -34.1602 | -9.9127
5.6037 | -34.1605 | -9.9232
5.6098 | -34.1608 | -9.9336
5.6159 | -34.1611 | -9.9439
5.622 | -34.1614 | -9.9544
5.6281 | -34.1617 | -9.965
5.6342 | -34.162 | -9.9757
5.6403 | -34.1623 | -9.9862
5.6464 | -34.1626 | -9.9968
5.6525 | -34.1629 | -10.0074
5.6586 | -34.1632 | -10.0175
5.6648 | -34.1636 | -10.028
5.6709 | -34.1639 | -10.0386
5.677 | -34.1642 | -10.0492
5.6831 | -34.1645 | -10.0597
5.6892 | -34.1648 | -10.0703
5.6953 | -34.1651 | -10.0808
5.7014 | -34.1654 | -10.0914
5.7075 | -34.1657 | -10.1018
5.7136 | -34.166 | -10.1123
5.7197 | -34.1663 | -10.1228
5.7258 | -34.1665 | -10.133
5.7319 | -34.1669 | -10.1436
5.738 | -34.1672 | -10.154
5.7441 | -34.1675 | -10.1645
5.7502 | -34.1679 | -10.1746
5.7563 | -34.1682 | -10.1851
5.7624 | -34.1685 | -10.1958
5.7685 | -34.1688 | -10.2063
5.7746 | -34.1691 | -10.217
5.7807 | -34.1693 | -10.2275
5.7868 | -34.1696 | -10.2382
5.7929 | -34.1699 | -10.2486
5.799 | -34.1702 | -10.2589
5.8052 | -34.1705 | -10.2692
5.8113 | -34.1708 | -10.2796
5.8174 | -34.1711 | -10.29
5.8235 | -34.1714 | -10.3004
5.8296 | -34.1717 | -10.311
5.8357 | -34.172 | -10.3215
5.8418 | -34.1723 | -10.332
5.8479 | -34.1726 | -10.3426
5.854 | -34.1729 | -10.3532
5.8601 | -34.1732 | -10.3635
5.8662 | -34.1735 | -10.3741
5.8723 | -34.1738 | -10.3843
5.8784 | -34.1741 | -10.3947
5.8845 | -34.1744 | -10.4051
5.8906 | -34.1746 | -10.4156
5.8967 | -34.1749 | -10.4258
5.9028 | -34.1753 | -10.4362
5.9089 | -34.1756 | -10.4465
5.915 | -34.1759 | -10.4567
5.9211 | -34.1762 | -10.4671
5.9272 | -34.1765 | -10.4773
5.9333 | -34.1768 | -10.4876
5.9394 | -34.1771 | -10.4981
5.9455 | -34.1774 | -10.5083
5.9517 | -34.1778 | -10.5185
5.9578 | -34.1781 | -10.5291
5.9639 | -34.1784 | -10.5392
5.97 | -34.1788 | -10.5496
5.9761 | -34.1791 | -10.56
5.9822 | -34.1794 | -10.5704
5.9883 | -34.1797 | -10.5808
5.9944 | -34.18 | -10.5913
6.0005 | -34.1803 | -10.6018
6.0066 | -34.1806 | -10.612
6.0127 | -34.181 | -10.6222
6.0188 | -34.1813 | -10.6328
6.0249 | -34.1816 | -10.6433
6.031 | -34.1819 | -10.6535
6.0371 | -34.1822 | -10.6639
6.0432 | -34.1825 | -10.6742
6.0493 | -34.1828 | -10.6843
6.0554 | -34.1831 | -10.6946
6.0615 | -34.1835 | -10.7049
6.0676 | -34.1838 | -10.7155
6.0737 | -34.1841 | -10.726
6.0798 | -34.1844 | -10.7363
6.0859 | -34.1847 | -10.7466
6.0921 | -34.185 | -10.7567
6.0982 | -34.1853 | -10.7671
6.1043 | -34.1857 | -10.7773
6.1104 | -34.186 | -10.7879
6.1165 | -34.1864 | -10.7977
6.1226 | -34.1867 | -10.808
6.1287 | -34.187 | -10.8183
6.1348 | -34.1873 | -10.8286
6.1409 | -34.1876 | -10.8389
6.147 | -34.1879 | -10.8491
6.1531 | -34.1882 | -10.8594
6.1592 | -34.1885 | -10.8696
6.1653 | -34.1888 | -10.8798
6.1714 | -34.1891 | -10.89
6.1775 | -34.1895 | -10.9004
6.1836 | -34.1898 | -10.9108
6.1897 | -34.1901 | -10.921
6.1958 | -34.1904 | -10.9316
6.2019 | -34.1907 | -10.942
6.208 | -34.1911 | -10.9523
6.2141 | -34.1914 | -10.9625
6.2202 | -34.1917 | -10.9729
6.2263 | -34.1921 | -10.983
6.2325 | -34.1923 | -10.993
6.2386 | -34.1927 | -11.003
6.2447 | -34.193 | -11.0135
6.2508 | -34.1933 | -11.0238
6.2569 | -34.1937 | -11.034
6.263 | -34.194 | -11.0442
6.2691 | -34.1943 | -11.0544
6.2752 | -34.1946 | -11.0647
6.2813 | -34.1949 | -11.075
6.2874 | -34.1952 | -11.0851
6.2935 | -34.1955 | -11.0953
6.2996 | -34.1958 | -11.1055
6.3057 | -34.1961 | -11.1158
6.3118 | -34.1964 | -11.1261
6.3179 | -34.1967 | -11.1363
6.324 | -34.197 | -11.1465
6.3301 | -34.1973 | -11.1567
6.3362 | -34.1977 | -11.1668
6.3423 | -34.198 | -11.1772
6.3484 | -34.1983 | -11.1875
6.3545 | -34.1986 | -11.1974
6.3606 | -34.1989 | -11.2074
6.3667 | -34.1991 | -11.2174
6.3728 | -34.1995 | -11.2277
6.379 | -34.1998 | -11.2379
6.3851 | -34.2001 | -11.2479
6.3912 | -34.2004 | -11.258
6.3973 | -34.2007 | -11.2682
6.4034 | -34.201 | -11.2782
6.4095 | -34.2013 | -11.2884
6.4156 | -34.2016 | -11.2986
6.4217 | -34.2019 | -11.3087
6.4278 | -34.2022 | -11.319
6.4339 | -34.2025 | -11.3291
6.44 | -34.2029 | -11.3392
6.4461 | -34.2032 | -11.3495
6.4522 | -34.2035 | -11.3595
6.4583 | -34.2038 | -11.3698
6.4644 | -34.204 | -11.3801
6.4705 | -34.2043 | -11.3898
6.4766 | -34.2047 | -11.4
6.4827 | -34.205 | -11.4102
6.4888 | -34.2053 | -11.4202
6.4949 | -34.2056 | -11.4301
6.501 | -34.2059 | -11.4403
6.5071 | -34.2062 | -11.4502
6.5132 | -34.2065 | -11.4604
6.5194 | -34.2068 | -11.4705
6.5255 | -34.2071 | -11.4805
6.5316 | -34.2074 | -11.4907
6.5377 | -34.2077 | -11.5008
6.5438 | -34.208 | -11.5107
6.5499 | -34.2083 | -11.5207
6.556 | -34.2086 | -11.5308
6.5621 | -34.2089 | -11.5411
6.5682 | -34.2092 | -11.5512
6.5743 | -34.2095 | -11.5613
6.5804 | -34.2098 | -11.5715
6.5865 | -34.2101 | -11.5816
6.5926 | -34.2105 | -11.5916
6.5987 | -34.2108 | -11.6014
6.6048 | -34.2111 | -11.6115
6.6109 | -34.2114 | -11.6216
6.617 | -34.2116 | -11.6318
6.6231 | -34.2119 | -11.6418
6.6292 | -34.2123 | -11.6518
6.6353 | -34.2126 | -11.6622
6.6414 | -34.2129 | -11.6722
6.6475 | -34.2132 | -11.6822
6.6536 | -34.2135 | -11.6922
6.6597 | -34.2138 | -11.702
6.6659 | -34.2141 | -11.712
6.672 | -34.2144 | -11.722
6.6781 | -34.2147 | -11.7321
6.6842 | -34.215 | -11.742
6.6903 | -34.2154 | -11.752
6.6964 | -34.2157 | -11.7621
6.7025 | -34.216 | -11.772
6.7086 | -34.2163 | -11.782
6.7147 | -34.2166 | -11.7921
6.7208 | -34.2169 | -11.8022
6.7269 | -34.2172 | -11.8122
6.733 | -34.2174 | -11.822
6.7391 | -34.2178 | -11.832
6.7452 | -34.2181 | -11.8419
6.7513 | -34.2184 | -11.8516
6.7574 | -34.2187 | -11.8617
6.7635 | -34.219 | -11.8715
6.7696 | -34.2193 | -11.8814
6.7757 | -34.2197 | -11.8914
6.7818 | -34.22 | -11.9014
6.7879 | -34.2203 | -11.9113
6.794 | -34.2206 | -11.9211
6.8001 | -34.2209 | -11.931
6.8063 | -34.2212 | -11.9409
6.8124 | -34.2215 | -11.951
6.8185 | -34.2218 | -11.9612
6.8246 | -34.222 | -11.9709
6.8307 | -34.2223 | -11.9808
6.8368 | -34.2227 | -11.991
6.8429 | -34.223 | -12.0009
6.849 | -34.2233 | -12.0109
6.8551 | -34.2237 | -12.0208
6.8612 | -34.2239 | -12.0307
6.8673 | -34.2242 | -12.0406
6.8734 | -34.2245 | -12.0507
6.8795 | -34.2249 | -12.0607
6.8856 | -34.2252 | -12.0706
6.8917 | -34.2255 | -12.0805
6.8978 | -34.2258 | -12.0902
6.9039 | -34.2261 | -12.0999
6.91 | -34.2264 | -12.1099
6.9161 | -34.2267 | -12.1196
6.9222 | -34.227 | -12.1295
6.9283 | -34.2274 | -12.1391
6.9344 | -34.2277 | -12.1491
6.9405 | -34.228 | -12.159
6.9466 | -34.2284 | -12.1691
6.9528 | -34.2287 | -12.1792
6.9589 | -34.229 | -12.1893
6.965 | -34.2293 | -12.1993
6.9711 | -34.2297 | -12.2093
6.9772 | -34.23 | -12.2192
6.9833 | -34.2303 | -12.2291
6.9894 | -34.2306 | -12.2393
6.9955 | -34.2309 | -12.2489
7.0016 | -34.2313 | -12.2588
7.0077 | -34.2316 | -12.2687
7.0138 | -34.2319 | -12.2785
7.0199 | -34.2322 | -12.2883
7.026 | -34.2325 | -12.2982
7.0321 | -34.2329 | -12.3078
7.0382 | -34.2332 | -12.3174
7.0443 | -34.2336 | -12.3275
7.0504 | -34.2339 | -12.3375
7.0565 | -34.2342 | -12.3474
7.0626 | -34.2345 | -12.357
7.0687 | -34.2348 | -12.3664
7.0748 | -34.2352 | -12.3761
7.0809 | -34.2355 | -12.3859
7.087 | -34.2358 | -12.3958
7.0932 | -34.2361 | -12.4057
7.0993 | -34.2364 | -12.4155
7.1054 | -34.2368 | -12.4253
7.1115 | -34.2371 | -12.4349
7.1176 | -34.2374 | -12.4447
7.1237 | -34.2377 | -12.4545
7.1298 | -34.2381 | -12.4643
7.1359 | -34.2384 | -12.4742
7.142 | -34.2387 | -12.4837
7.1481 | -34.239 | -12.4935
7.1542 | -34.2393 | -12.5034
7.1603 | -34.2397 | -12.5132
7.1664 | -34.24 | -12.523
7.1725 | -34.2404 | -12.5329
7.1786 | -34.2407 | -12.5427
7.1847 | -34.241 | -12.5525
7.1908 | -34.2413 | -12.5625
7.1969 | -34.2416 | -12.5725
7.203 | -34.242 | -12.5821
7.2091 | -34.2423 | -12.5919
7.2152 | -34.2427 | -12.6017
7.2213 | -34.243 | -12.6113
7.2274 | -34.2433 | -12.6211
7.2335 | -34.2436 | -12.6308
7.2397 | -34.244 | -12.6406
7.2458 | -34.2444 | -12.6501
7.2519 | -34.2447 | -12.6599
7.258 | -34.245 | -12.6698
7.2641 | -34.2453 | -12.6797
7.2702 | -34.2456 | -12.6896
7.2763 | -34.246 | -12.6994
7.2824 | -34.2464 | -12.7094
7.2885 | -34.2467 | -12.7193
7.2946 | -34.247 | -12.7291
7.3007 | -34.2473 | -12.739
7.3068 | -34.2476 | -12.749
7.3129 | -34.248 | -12.759
7.319 | -34.2483 | -12.7688
7.3251 | -34.2486 | -12.7784
7.3312 | -34.249 | -12.7882
7.3373 | -34.2493 | -12.7978
7.3434 | -34.2497 | -12.8075
7.3495 | -34.2501 | -12.8172
7.3556 | -34.2504 | -12.8272
7.3617 | -34.2507 | -12.8372
7.3678 | -34.2511 | -12.8469
7.3739 | -34.2514 | -12.8569
7.3801 | -34.2517 | -12.8669
7.3862 | -34.2521 | -12.8767
7.3923 | -34.2524 | -12.8867
7.3984 | -34.2527 | -12.8962
7.4045 | -34.2531 | -12.906
7.4106 | -34.2535 | -12.9157
7.4167 | -34.2538 | -12.9256
7.4228 | -34.2542 | -12.9355
7.4289 | -34.2545 | -12.9453
7.435 | -34.2549 | -12.9549
7.4411 | -34.2552 | -12.9649
7.4472 | -34.2555 | -12.9747
7.4533 | -34.2559 | -12.9842
7.4594 | -34.2562 | -12.9943
7.4655 | -34.2566 | -13.0039
7.4716 | -34.2569 | -13.0138
7.4777 | -34.2573 | -13.0236
7.4838 | -34.2576 | -13.0332
7.4899 | -34.2579 | -13.0429
7.496 | -34.2583 | -13.0525
7.5021 | -34.2586 | -13.0623
7.5082 | -34.259 | -13.0721
7.5143 | -34.2593 | -13.0819
7.5204 | -34.2597 | -13.0919
7.5266 | -34.26 | -13.1014
7.5327 | -34.2604 | -13.1111
7.5388 | -34.2607 | -13.1213
7.5449 | -34.2611 | -13.1312
7.551 | -34.2614 | -13.1409
7.5571 | -34.2618 | -13.1508
7.5632 | -34.2621 | -13.1606
7.5693 | -34.2625 | -13.1703
7.5754 | -34.2628 | -13.1799
7.5815 | -34.2632 | -13.1898
7.5876 | -34.2635 | -13.1998
7.5937 | -34.2639 | -13.2098
7.5998 | -34.2643 | -13.2194
7.6059 | -34.2647 | -13.2294
7.612 | -34.265 | -13.2392
7.6181 | -34.2654 | -13.2489
7.6242 | -34.2657 | -13.2591
7.6303 | -34.2661 | -13.2691
7.6364 | -34.2664 | -13.279
7.6425 | -34.2668 | -13.2889
7.6486 | -34.2671 | -13.2988
7.6547 | -34.2675 | -13.3085
7.6608 | -34.2678 | -13.3182
7.667 | -34.2681 | -13.3282
7.6731 | -34.2685 | -13.3382
7.6792 | -34.2689 | -13.3482
7.6853 | -34.2692 | -13.3581
7.6914 | -34.2696 | -13.368
7.6975 | -34.27 | -13.3779
7.7036 | -34.2703 | -13.3878
7.7097 | -34.2707 | -13.3977
7.7158 | -34.2711 | -13.4076
7.7219 | -34.2714 | -13.4172
7.728 | -34.2718 | -13.4272
7.7341 | -34.2721 | -13.4372
7.7402 | -34.2725 | -13.4471
7.7463 | -34.2728 | -13.457
7.7524 | -34.2732 | -13.4672
7.7585 | -34.2735 | -13.477
7.7646 | -34.2739 | -13.487
7.7707 | -34.2742 | -13.4972
7.7768 | -34.2746 | -13.5072
7.7829 | -34.2749 | -13.5175
7.789 | -34.2753 | -13.5277
7.7951 | -34.2756 | -13.5374
7.8012 | -34.276 | -13.5472
7.8073 | -34.2764 | -13.5573
7.8135 | -34.2767 | -13.5671
7.8196 | -34.2771 | -13.5772
7.8257 | -34.2775 | -13.5873
7.8318 | -34.2778 | -13.5974
7.8379 | -34.2782 | -13.6071
7.844 | -34.2785 | -13.617
7.8501 | -34.2789 | -13.6272
7.8562 | -34.2792 | -13.6373
7.8623 | -34.2796 | -13.6471
7.8684 | -34.28 | -13.6572
7.8745 | -34.2803 | -13.6672
7.8806 | -34.2807 | -13.6774
7.8867 | -34.281 | -13.687
7.8928 | -34.2814 | -13.6967
7.8989 | -34.2818 | -13.7067
7.905 | -34.2821 | -13.7168
7.9111 | -34.2824 | -13.727
7.9172 | -34.2828 | -13.7369
7.9233 | -34.2831 | -13.747
7.9294 | -34.2835 | -13.757
7.9355 | -34.2839 | -13.7668
7.9416 | -34.2842 | -13.7768
7.9477 | -34.2845 | -13.7867
7.9539 | -34.2849 | -13.7967
7.96 | -34.2853 | -13.8068
7.9661 | -34.2856 | -13.8168
7.9722 | -34.286 | -13.8265
7.9783 | -34.2864 | -13.8368
7.9844 | -34.2868 | -13.8466
7.9905 | -34.2871 | -13.8568
7.9966 | -34.2875 | -13.8669
8.0027 | -34.2878 | -13.8769
8.0088 | -34.2882 | -13.8868
8.0149 | -34.2885 | -13.8969
8.021 | -34.2889 | -13.907
8.0271 | -34.2893 | -13.9172
8.0332 | -34.2896 | -13.9272
8.0393 | -34.29 | -13.9372
8.0454 | -34.2904 | -13.9474
8.0515 | -34.2908 | -13.9573
8.0576 | -34.2911 | -13.9676
8.0637 | -34.2914 | -13.9776
8.0698 | -34.2918 | -13.9874
8.0759 | -34.2921 | -13.9974
8.082 | -34.2925 | -14.0073
8.0881 | -34.2928 | -14.0175
8.0942 | -34.2932 | -14.0275
8.1004 | -34.2936 | -14.0376
8.1065 | -34.294 | -14.0475
8.1126 | -34.2943 | -14.0575
8.1187 | -34.2947 | -14.0678
8.1248 | -34.295 | -14.0778
8.1309 | -34.2954 | -14.0877
8.137 | -34.2957 | -14.0975
8.1431 | -34.2961 | -14.1077
8.1492 | -34.2964 | -14.1178
8.1553 | -34.2968 | -14.1277
8.1614 | -34.2972 | -14.1379
8.1675 | -34.2975 | -14.1478
8.1736 | -34.2979 | -14.1581
8.1797 | -34.2982 | -14.1681
8.1858 | -34.2986 | -14.1782
8.1919 | -34.2989 | -14.1883
8.198 | -34.2993 | -14.1982
8.2041 | -34.2997 | -14.2081
8.2102 | -34.3001 | -14.218
8.2163 | -34.3004 | -14.228
8.2224 | -34.3008 | -14.238
8.2285 | -34.3011 | -14.248
8.2346 | -34.3015 | -14.258
8.2408 | -34.3019 | -14.2679
8.2469 | -34.3022 | -14.278
8.253 | -34.3026 | -14.288
8.2591 | -34.303 | -14.298
8.2652 | -34.3033 | -14.308
8.2713 | -34.3037 | -14.318
8.2774 | -34.3041 | -14.3281
8.2835 | -34.3044 | -14.3383
8.2896 | -34.3048 | -14.3482
8.2957 | -34.3052 | -14.3583
8.3018 | -34.3055 | -14.3683
8.3079 | -34.3059 | -14.3782
8.314 | -34.3063 | -14.3881
8.3201 | -34.3067 | -14.398
8.3262 | -34.307 | -14.408
8.3323 | -34.3074 | -14.418
8.3384 | -34.3077 | -14.4282
8.3445 | -34.3081 | -14.4382
8.3506 | -34.3085 | -14.4482
8.3567 | -34.3088 | -14.4583
8.3628 | -34.3092 | -14.4684
8.3689 | -34.3095 | -14.4784
8.375 | -34.3099 | -14.4885
8.3812 | -34.3102 | -14.4985
8.3873 | -34.3106 | -14.5085
8.3934 | -34.3109 | -14.5186
8.3995 | -34.3113 | -14.5286
8.4056 | -34.3116 | -14.5385
8.4117 | -34.312 | -14.5486
8.4178 | -34.3123 | -14.5586
8.4239 | -34.3127 | -14.5686
8.43 | -34.3131 | -14.5786
8.4361 | -34.3134 | -14.5887
8.4422 | -34.3138 | -14.5986
8.4483 | -34.3141 | -14.6086
8.4544 | -34.3145 | -14.6186
8.4605 | -34.3148 | -14.6285
8.4666 | -34.3152 | -14.6387
8.4727 | -34.3155 | -14.6486
8.4788 | -34.3159 | -14.6585
8.4849 | -34.3162 | -14.6686
8.491 | -34.3166 | -14.6786
8.4971 | -34.3169 | -14.6884
8.5032 | -34.3173 | -14.6984
8.5093 | -34.3176 | -14.7083
8.5154 | -34.318 | -14.7182
8.5215 | -34.3183 | -14.7283
8.5277 | -34.3187 | -14.7381
8.5338 | -34.319 | -14.748
8.5399 | -34.3194 | -14.7581
8.546 | -34.3197 | -14.7679
8.5521 | -34.3201 | -14.778
8.5582 | -34.3204 | -14.7878
8.5643 | -34.3208 | -14.7976
8.5704 | -34.3212 | -14.8074
8.5765 | -34.3215 | -14.8174
8.5826 | -34.3219 | -14.8272
8.5887 | -34.3223 | -14.8371
8.5948 | -34.3226 | -14.8471
8.6009 | -34.323 | -14.8569
8.607 | -34.3233 | -14.867
8.6131 | -34.3237 | -14.8771
8.6192 | -34.3241 | -14.887
8.6253 | -34.3244 | -14.8968
8.6314 | -34.3248 | -14.9066
8.6375 | -34.3251 | -14.9165
8.6436 | -34.3255 | -14.9264
8.6497 | -34.3258 | -14.9361
8.6558 | -34.3262 | -14.946
8.6619 | -34.3265 | -14.956
8.6681 | -34.3268 | -14.966
8.6742 | -34.3272 | -14.9759
8.6803 | -34.3275 | -14.9859
8.6864 | -34.3279 | -14.9957
8.6925 | -34.3283 | -15.0054
8.6986 | -34.3286 | -15.0155
8.7047 | -34.329 | -15.0255
8.7108 | -34.3293 | -15.0354
8.7169 | -34.3296 | -15.0453
8.723 | -34.33 | -15.0548
8.7291 | -34.3304 | -15.0647
8.7352 | -34.3307 | -15.0748
8.7413 | -34.331 | -15.0846
8.7474 | -34.3314 | -15.0943
8.7535 | -34.3317 | -15.1043
8.7596 | -34.3321 | -15.1141
8.7657 | -34.3324 | -15.124
8.7718 | -34.3328 | -15.1338
8.7779 | -34.3331 | -15.1439
8.784 | -34.3335 | -15.1538
8.7901 | -34.3338 | -15.1637
8.7962 | -34.3342 | -15.1733
8.8023 | -34.3345 | -15.1831
8.8084 | -34.3349 | -15.1931
8.8146 | -34.3352 | -15.2029
8.8207 | -34.3356 | -15.2125
8.8268 | -34.336 | -15.2221
8.8329 | -34.3363 | -15.2318
8.839 | -34.3367 | -15.2417
8.8451 | -34.337 | -15.2514
8.8512 | -34.3373 | -15.2613
8.8573 | -34.3377 | -15.2711
8.8634 | -34.338 | -15.2808
8.8695 | -34.3384 | -15.2907
8.8756 | -34.3387 | -15.3004
8.8817 | -34.3391 | -15.3102
8.8878 | -34.3394 | -15.3202
8.8939 | -34.3397 | -15.33
8.9 | -34.3401 | -15.3397
8.9061 | -34.3405 | -15.3495
8.9122 | -34.3408 | -15.3592
8.9183 | -34.3411 | -15.3691
8.9244 | -34.3415 | -15.3791
8.9305 | -34.3419 | -15.3887
8.9366 | -34.3423 | -15.3985
8.9427 | -34.3426 | -15.4084
8.9488 | -34.343 | -15.4181
8.955 | -34.3433 | -15.4278
8.9611 | -34.3437 | -15.4378
8.9672 | -34.344 | -15.4476
8.9733 | -34.3444 | -15.4571
8.9794 | -34.3448 | -15.4668
8.9855 | -34.3451 | -15.4766
8.9916 | -34.3455 | -15.4865
8.9977 | -34.3458 | -15.4962
9.0038 | -34.3462 | -15.5059
9.0099 | -34.3466 | -15.5157
9.016 | -34.3469 | -15.5254
9.0221 | -34.3473 | -15.535
9.0282 | -34.3476 | -15.5446
9.0343 | -34.348 | -15.5544
9.0404 | -34.3484 | -15.5641
9.0465 | -34.3487 | -15.5739
9.0526 | -34.3491 | -15.5836
9.0587 | -34.3495 | -15.5934
9.0648 | -34.3499 | -15.6031
9.0709 | -34.3502 | -15.6127
9.077 | -34.3506 | -15.6224
9.0831 | -34.3509 | -15.6323
9.0892 | -34.3513 | -15.642
9.0953 | -34.3517 | -15.6519
9.1015 | -34.352 | -15.6616
9.1076 | -34.3524 | -15.6711
9.1137 | -34.3528 | -15.6807
9.1198 | -34.3531 | -15.6905
9.1259 | -34.3535 | -15.7002
9.132 | -34.3539 | -15.71
9.1381 | -34.3543 | -15.7197
9.1442 | -34.3547 | -15.729
9.1503 | -34.355 | -15.7388
9.1564 | -34.3554 | -15.7485
9.1625 | -34.3557 | -15.7581
9.1686 | -34.3561 | -15.7678
9.1747 | -34.3565 | -15.7775
9.1808 | -34.3569 | -15.7871
9.1869 | -34.3572 | -15.7966
9.193 | -34.3576 | -15.8063
9.1991 | -34.3579 | -15.816
9.2052 | -34.3583 | -15.8257
9.2113 | -34.3587 | -15.8354
9.2174 | -34.3591 | -15.8451
9.2235 | -34.3595 | -15.8547
9.2296 | -34.3599 | -15.8644
9.2357 | -34.3602 | -15.8742
9.2419 | -34.3606 | -15.8839
9.248 | -34.361 | -15.8936
9.2541 | -34.3614 | -15.9034
9.2602 | -34.3618 | -15.9131
9.2663 | -34.3621 | -15.9225
9.2724 | -34.3625 | -15.9321
9.2785 | -34.3628 | -15.9416
9.2846 | -34.3633 | -15.9511
9.2907 | -34.3636 | -15.9607
9.2968 | -34.364 | -15.9705
9.3029 | -34.3643 | -15.9803
9.309 | -34.3647 | -15.9899
9.3151 | -34.3651 | -15.9994
9.3212 | -34.3655 | -16.0088
9.3273 | -34.3658 | -16.0183
9.3334 | -34.3662 | -16.0277
9.3395 | -34.3666 | -16.0371
9.3456 | -34.3669 | -16.0465
9.3517 | -34.3673 | -16.0562
9.3578 | -34.3677 | -16.0657
9.3639 | -34.3681 | -16.0752
9.37 | -34.3684 | -16.085
9.3761 | -34.3688 | -16.0947
9.3822 | -34.3692 | -16.1044
9.3884 | -34.3696 | -16.1135
9.3945 | -34.37 | -16.1233
9.4006 | -34.3704 | -16.1331
9.4067 | -34.3708 | -16.1428
9.4128 | -34.3711 | -16.1525
9.4189 | -34.3715 | -16.1624
9.425 | -34.3719 | -16.1719
9.4311 | -34.3723 | -16.1816
9.4372 | -34.3727 | -16.1913
9.4433 | -34.3731 | -16.2008
9.4494 | -34.3735 | -16.2102
9.4555 | -34.3739 | -16.2199
9.4616 | -34.3743 | -16.2294
9.4677 | -34.3746 | -16.239
9.4738 | -34.375 | -16.2485
9.4799 | -34.3754 | -16.2581
9.486 | -34.3758 | -16.2676
9.4921 | -34.3761 | -16.2772
9.4982 | -34.3766 | -16.2865
9.5043 | -34.377 | -16.296
9.5104 | -34.3773 | -16.3054
9.5165 | -34.3777 | -16.3148
9.5226 | -34.3781 | -16.3246
9.5288 | -34.3785 | -16.3337
9.5349 | -34.3788 | -16.3432
9.541 | -34.3793 | -16.3526
9.5471 | -34.3796 | -16.3621
9.5532 | -34.38 | -16.3718
9.5593 | -34.3804 | -16.3814
9.5654 | -34.3808 | -16.3907
9.5715 | -34.3812 | -16.4002
9.5776 | -34.3815 | -16.41
9.5837 | -34.3819 | -16.4194
9.5898 | -34.3822 | -16.429
9.5959 | -34.3827 | -16.4385
9.602 | -34.383 | -16.4476
9.6081 | -34.3834 | -16.457
9.6142 | -34.3838 | -16.4666
9.6203 | -34.3842 | -16.476
9.6264 | -34.3846 | -16.4853
9.6325 | -34.3849 | -16.495
9.6386 | -34.3853 | -16.5043
9.6447 | -34.3857 | -16.5138
9.6508 | -34.3861 | -16.5232
9.6569 | -34.3864 | -16.5325
9.663 | -34.3869 | -16.5422
9.6691 | -34.3873 | -16.5519
9.6753 | -34.3877 | -16.5612
9.6814 | -34.3881 | -16.5705
9.6875 | -34.3885 | -16.5798
9.6936 | -34.3888 | -16.5891
9.6997 | -34.3892 | -16.5987
9.7058 | -34.3897 | -16.6082
9.7119 | -34.3901 | -16.6178
9.718 | -34.3905 | -16.6275
9.7241 | -34.3909 | -16.6368
9.7302 | -34.3912 | -16.6464
9.7363 | -34.3917 | -16.6557
9.7424 | -34.3921 | -16.665
9.7485 | -34.3924 | -16.6745
9.7546 | -34.3928 | -16.6837
9.7607 | -34.3932 | -16.6934
9.7668 | -34.3936 | -16.7029
9.7729 | -34.394 | -16.7123
9.779 | -34.3945 | -16.7215
9.7851 | -34.3948 | -16.7311
9.7912 | -34.3953 | -16.7405
9.7973 | -34.3957 | -16.7503
9.8034 | -34.3961 | -16.7594
9.8095 | -34.3965 | -16.7691
9.8157 | -34.397 | -16.7786
9.8218 | -34.3974 | -16.7882
9.8279 | -34.3978 | -16.7974
9.834 | -34.3981 | -16.807
9.8401 | -34.3985 | -16.8164
9.8462 | -34.399 | -16.8254
9.8523 | -34.3993 | -16.8348
9.8584 | -34.3998 | -16.8441
9.8645 | -34.4001 | -16.8537
9.8706 | -34.4005 | -16.8632
9.8767 | -34.4009 | -16.8726
9.8828 | -34.4013 | -16.8818
9.8889 | -34.4017 | -16.8909
9.895 | -34.4021 | -16.9004
9.9011 | -34.4025 | -16.9098
9.9072 | -34.4029 | -16.9192
9.9133 | -34.4033 | -16.9286
9.9194 | -34.4037 | -16.938
9.9255 | -34.4041 | -16.9472
9.9316 | -34.4045 | -16.9567
9.9377 | -34.405 | -16.9661
9.9438 | -34.4053 | -16.9757
9.9499 | -34.4058 | -16.9854
9.956 | -34.4062 | -16.9948
9.9622 | -34.4066 | -17.0042
9.9683 | -34.407 | -17.0136
9.9744 | -34.4073 | -17.0233
9.9805 | -34.4078 | -17.0325
9.9866 | -34.4082 | -17.0419
9.9927 | -34.4086 | -17.0511
9.9988 | -34.409 | -17.0606
10.0049 | -34.4094 | -17.0699
10.011 | -34.4098 | -17.0796
10.0171 | -34.4102 | -17.0888
10.0232 | -34.4107 | -17.0982
10.0293 | -34.4111 | -17.1076
10.0354 | -34.4115 | -17.1169
10.0415 | -34.4119 | -17.1263
10.0476 | -34.4123 | -17.1359
10.0537 | -34.4127 | -17.1454
10.0598 | -34.4132 | -17.1548
10.0659 | -34.4135 | -17.1643
10.072 | -34.4139 | -17.1734
10.0781 | -34.4143 | -17.1826
10.0842 | -34.4148 | -17.1919
10.0903 | -34.4152 | -17.2012
10.0964 | -34.4156 | -17.2102
10.1026 | -34.416 | -17.2196
10.1087 | -34.4164 | -17.2291
10.1148 | -34.4168 | -17.2386
10.1209 | -34.4172 | -17.2478
10.127 | -34.4176 | -17.257
10.1331 | -34.418 | -17.2662
10.1392 | -34.4184 | -17.2753
10.1453 | -34.4188 | -17.2846
10.1514 | -34.4192 | -17.2942
10.1575 | -34.4196 | -17.3038
10.1636 | -34.42 | -17.3128
10.1697 | -34.4204 | -17.3222
10.1758 | -34.4208 | -17.3314
10.1819 | -34.4212 | -17.3405
10.188 | -34.4216 | -17.3499
10.1941 | -34.4221 | -17.3593
10.2002 | -34.4225 | -17.369
10.2063 | -34.4229 | -17.3782
10.2124 | -34.4233 | -17.3875
10.2185 | -34.4237 | -17.3967
10.2246 | -34.4241 | -17.4061
10.2307 | -34.4245 | -17.4154
10.2368 | -34.4249 | -17.4244
10.2429 | -34.4253 | -17.4336
10.2491 | -34.4257 | -17.4431
10.2552 | -34.4261 | -17.4523
10.2613 | -34.4266 | -17.4615
10.2674 | -34.427 | -17.4708
10.2735 | -34.4274 | -17.4803
10.2796 | -34.4278 | -17.4897
10.2857 | -34.4283 | -17.499
10.2918 | -34.4287 | -17.5085
10.2979 | -34.4291 | -17.5178
10.304 | -34.4296 | -17.5267
10.3101 | -34.43 | -17.5361
10.3162 | -34.4304 | -17.5454
10.3223 | -34.4308 | -17.5546
10.3284 | -34.4312 | -17.5639
10.3345 | -34.4316 | -17.5734
10.3406 | -34.4321 | -17.5824
10.3467 | -34.4325 | -17.5918
10.3528 | -34.4329 | -17.601
10.3589 | -34.4333 | -17.6102
10.365 | -34.4337 | -17.6195
10.3711 | -34.4342 | -17.6289
10.3772 | -34.4347 | -17.6377
10.3833 | -34.4352 | -17.6465
10.3895 | -34.4356 | -17.6555
10.3956 | -34.436 | -17.6646
10.4017 | -34.4365 | -17.6738
10.4078 | -34.4368 | -17.683
10.4139 | -34.4373 | -17.6925
10.42 | -34.4378 | -17.7019
10.4261 | -34.4382 | -17.7111
10.4322 | -34.4386 | -17.7203
10.4383 | -34.439 | -17.7297
10.4444 | -34.4395 | -17.7387
10.4505 | -34.4399 | -17.7481
10.4566 | -34.4403 | -17.757
10.4627 | -34.4407 | -17.7663
10.4688 | -34.4411 | -17.7752
10.4749 | -34.4415 | -17.7843
10.481 | -34.4419 | -17.7933
10.4871 | -34.4423 | -17.8023
10.4932 | -34.4428 | -17.8113
10.4993 | -34.4432 | -17.8205
10.5054 | -34.4436 | -17.8301
10.5115 | -34.444 | -17.8391
10.5176 | -34.4444 | -17.8477
10.5237 | -34.4448 | -17.8568
10.5298 | -34.4453 | -17.866
10.536 | -34.4457 | -17.8751
10.5421 | -34.4461 | -17.8843
10.5482 | -34.4466 | -17.8935
10.5543 | -34.447 | -17.9022
10.5604 | -34.4474 | -17.9116
10.5665 | -34.4479 | -17.9209
10.5726 | -34.4483 | -17.9304
10.5787 | -34.4487 | -17.9393
10.5848 | -34.4492 | -17.9481
10.5909 | -34.4496 | -17.9572
10.597 | -34.4501 | -17.9667
10.6031 | -34.4505 | -17.9758
10.6092 | -34.4509 | -17.9847
10.6153 | -34.4514 | -17.9938
10.6214 | -34.4518 | -18.003
10.6275 | -34.4522 | -18.0122
10.6336 | -34.4527 | -18.0212
10.6397 | -34.4531 | -18.0308
10.6458 | -34.4535 | -18.0401
10.6519 | -34.454 | -18.0493
10.658 | -34.4544 | -18.0578
10.6641 | -34.4548 | -18.0672
10.6702 | -34.4552 | -18.0763
10.6764 | -34.4557 | -18.0853
10.6825 | -34.4561 | -18.0943
10.6886 | -34.4565 | -18.1034
10.6947 | -34.4569 | -18.1122
10.7008 | -34.4573 | -18.1215
10.7069 | -34.4578 | -18.1306
10.713 | -34.4582 | -18.1398
10.7191 | -34.4586 | -18.1489
10.7252 | -34.4591 | -18.1579
10.7313 | -34.4594 | -18.167
10.7374 | -34.4598 | -18.1759
10.7435 | -34.4603 | -18.1846
10.7496 | -34.4607 | -18.1935
10.7557 | -34.4611 | -18.2023
10.7618 | -34.4615 | -18.2111
10.7679 | -34.462 | -18.2202
10.774 | -34.4624 | -18.2291
10.7801 | -34.4629 | -18.2379
10.7862 | -34.4633 | -18.2469
10.7923 | -34.4637 | -18.2561
10.7984 | -34.4642 | -18.2647
10.8045 | -34.4646 | -18.2733
10.8106 | -34.465 | -18.2828
10.8168 | -34.4654 | -18.2917
10.8229 | -34.4659 | -18.3004
10.829 | -34.4663 | -18.3096
10.8351 | -34.4669 | -18.3184
10.8412 | -34.4674 | -18.3272
10.8473 | -34.4679 | -18.336
10.8534 | -34.4683 | -18.3449
10.8595 | -34.4688 | -18.3542
10.8656 | -34.4693 | -18.3628
10.8717 | -34.4697 | -18.3717
10.8778 | -34.4702 | -18.381
10.8839 | -34.4706 | -18.3899
10.89 | -34.471 | -18.3988
10.8961 | -34.4715 | -18.4077
10.9022 | -34.472 | -18.4166
10.9083 | -34.4725 | -18.4256
10.9144 | -34.4729 | -18.4343
10.9205 | -34.4733 | -18.4436
10.9266 | -34.4738 | -18.453
10.9327 | -34.4742 | -18.462
10.9388 | -34.4748 | -18.4706
10.9449 | -34.4752 | -18.4797
10.951 | -34.4757 | -18.4887
10.9571 | -34.4761 | -18.4973
10.9633 | -34.4765 | -18.5065
10.9694 | -34.4769 | -18.5156
10.9755 | -34.4774 | -18.5247
10.9816 | -34.4778 | -18.5341
10.9877 | -34.4783 | -18.5429
10.9938 | -34.4787 | -18.5519
10.9999 | -34.4792 | -18.5608
11.006 | -34.4797 | -18.569
11.0121 | -34.4801 | -18.5776
11.0182 | -34.4806 | -18.5866
11.0243 | -34.481 | -18.5957
11.0304 | -34.4814 | -18.6049
11.0365 | -34.4819 | -18.6142
11.0426 | -34.4824 | -18.6229
11.0487 | -34.4828 | -18.6321
11.0548 | -34.4833 | -18.6406
11.0609 | -34.4838 | -18.6501
11.067 | -34.4842 | -18.6592
11.0731 | -34.4846 | -18.6682
11.0792 | -34.4851 | -18.6771
11.0853 | -34.4856 | -18.6866
11.0914 | -34.486 | -18.6961
11.0975 | -34.4865 | -18.7052
11.1037 | -34.4869 | -18.7142
11.1098 | -34.4874 | -18.7237
11.1159 | -34.4878 | -18.7325
11.122 | -34.4882 | -18.741
11.1281 | -34.4886 | -18.7499
11.1342 | -34.4891 | -18.7589
11.1403 | -34.4896 | -18.768
11.1464 | -34.49 | -18.7771
11.1525 | -34.4904 | -18.7862
11.1586 | -34.4909 | -18.7954
11.1647 | -34.4914 | -18.8044
11.1708 | -34.4918 | -18.8131
11.1769 | -34.4923 | -18.822
11.183 | -34.4927 | -18.831
11.1891 | -34.4932 | -18.8399
11.1952 | -34.4937 | -18.8485
11.2013 | -34.4941 | -18.8574
11.2074 | -34.4946 | -18.8662
11.2135 | -34.495 | -18.8752
11.2196 | -34.4955 | -18.8844
11.2257 | -34.496 | -18.8936
11.2318 | -34.4966 | -18.9025
11.2379 | -34.497 | -18.9117
11.244 | -34.4976 | -18.9207
11.2502 | -34.498 | -18.9294
11.2563 | -34.4985 | -18.938
11.2624 | -34.4989 | -18.9465
11.2685 | -34.4994 | -18.9551
11.2746 | -34.4998 | -18.964
11.2807 | -34.5003 | -18.9731
11.2868 | -34.5008 | -18.9819
11.2929 | -34.5013 | -18.9905
11.299 | -34.5017 | -18.9995
11.3051 | -34.5022 | -19.009
11.3112 | -34.5027 | -19.0183
11.3173 | -34.5031 | -19.0269
11.3234 | -34.5035 | -19.0351
11.3295 | -34.504 | -19.0437
11.3356 | -34.5045 | -19.0524
11.3417 | -34.505 | -19.061
11.3478 | -34.5054 | -19.0702
11.3539 | -34.5058 | -19.0793
11.36 | -34.5063 | -19.0884
11.3661 | -34.5067 | -19.097
11.3722 | -34.5072 | -19.106
11.3783 | -34.5077 | -19.1151
11.3844 | -34.5081 | -19.124
11.3906 | -34.5086 | -19.1331
11.3967 | -34.5091 | -19.1415
11.4028 | -34.5095 | -19.1503
11.4089 | -34.51 | -19.159
11.415 | -34.5104 | -19.1682
11.4211 | -34.5109 | -19.1771
11.4272 | -34.5113 | -19.1859
11.4333 | -34.5118 | -19.1943
11.4394 | -34.5122 | -19.2031
11.4455 | -34.5126 | -19.2117
11.4516 | -34.5131 | -19.2205
11.4577 | -34.5136 | -19.2295
11.4638 | -34.514 | -19.2385
11.4699 | -34.5145 | -19.2472
11.476 | -34.515 | -19.2565
11.4821 | -34.5155 | -19.2652
11.4882 | -34.516 | -19.274
11.4943 | -34.5164 | -19.2831
11.5004 | -34.5169 | -19.2919
11.5065 | -34.5174 | -19.3004
11.5126 | -34.5178 | -19.3095
11.5187 | -34.5183 | -19.3183
11.5248 | -34.5187 | -19.3273
11.5309 | -34.5192 | -19.3361
11.5371 | -34.5197 | -19.3452
11.5432 | -34.5202 | -19.354
11.5493 | -34.5206 | -19.363
11.5554 | -34.5211 | -19.3718
11.5615 | -34.5216 | -19.3809
11.5676 | -34.5221 | -19.3897
11.5737 | -34.5225 | -19.3987
11.5798 | -34.523 | -19.4073
11.5859 | -34.5235 | -19.4161
11.592 | -34.524 | -19.4249
11.5981 | -34.5244 | -19.4339
11.6042 | -34.5249 | -19.443
11.6103 | -34.5253 | -19.4517
11.6164 | -34.5258 | -19.4603
11.6225 | -34.5262 | -19.4688
11.6286 | -34.5267 | -19.4775
11.6347 | -34.5272 | -19.4865
11.6408 | -34.5276 | -19.4955
11.6469 | -34.5281 | -19.5044
11.653 | -34.5286 | -19.5131
11.6591 | -34.529 | -19.5221
11.6652 | -34.5295 | -19.5311
11.6713 | -34.5299 | -19.5398
11.6775 | -34.5304 | -19.5487
11.6836 | -34.5309 | -19.5577
11.6897 | -34.5314 | -19.5665
11.6958 | -34.5318 | -19.5752
11.7019 | -34.5323 | -19.5842
11.708 | -34.5327 | -19.593
11.7141 | -34.5333 | -19.6015
11.7202 | -34.5337 | -19.6105
11.7263 | -34.5342 | -19.6194
11.7324 | -34.5346 | -19.6282
11.7385 | -34.5351 | -19.6373
11.7446 | -34.5355 | -19.6459
11.7507 | -34.536 | -19.6545
11.7568 | -34.5365 | -19.6635
11.7629 | -34.537 | -19.6724
11.769 | -34.5375 | -19.6812
11.7751 | -34.5379 | -19.6901
11.7812 | -34.5384 | -19.6989
11.7873 | -34.5389 | -19.7076
11.7934 | -34.5393 | -19.7166
11.7995 | -34.5398 | -19.7253
11.8056 | -34.5402 | -19.7341
11.8117 | -34.5407 | -19.743
11.8178 | -34.5412 | -19.7517
11.824 | -34.5416 | -19.7607
11.8301 | -34.5421 | -19.7695
11.8362 | -34.5426 | -19.7782
11.8423 | -34.543 | -19.7871
11.8484 | -34.5435 | -19.796
11.8545 | -34.5439 | -19.8049
11.8606 | -34.5444 | -19.8137
11.8667 | -34.5449 | -19.8225
11.8728 | -34.5453 | -19.8313
11.8789 | -34.5458 | -19.84
11.885 | -34.5463 | -19.8489
11.8911 | -34.5467 | -19.8577
11.8972 | -34.5472 | -19.8665
11.9033 | -34.5477 | -19.8754
11.9094 | -34.5481 | -19.8842
11.9155 | -34.5486 | -19.8929
11.9216 | -34.5491 | -19.9018
11.9277 | -34.5496 | -19.9106
11.9338 | -34.55 | -19.9195
11.9399 | -34.5505 | -19.9283
11.946 | -34.551 | -19.9372
11.9521 | -34.5514 | -19.9459
11.9582 | -34.5519 | -19.9549
11.9644 | -34.5524 | -19.9636
11.9705 | -34.5529 | -19.9724
11.9766 | -34.5534 | -19.9812
11.9827 | -34.5538 | -19.9899
11.9888 | -34.5543 | -19.9987
11.9949 | -34.5548 | -20.0075
12.001 | -34.5552 | -20.0166
12.0071 | -34.5557 | -20.0254
12.0132 | -34.5562 | -20.034
12.0193 | -34.5566 | -20.043
12.0254 | -34.5571 | -20.0517
12.0315 | -34.5575 | -20.0605
12.0376 | -34.558 | -20.0692
12.0437 | -34.5584 | -20.078
12.0498 | -34.5589 | -20.087
12.0559 | -34.5594 | -20.0957
12.062 | -34.5599 | -20.1044
12.0681 | -34.5603 | -20.1132
12.0742 | -34.5608 | -20.122
12.0803 | -34.5612 | -20.1309
12.0864 | -34.5617 | -20.1398
12.0925 | -34.5622 | -20.1487
12.0986 | -34.5627 | -20.1576
12.1047 | -34.5631 | -20.1663
12.1109 | -34.5636 | -20.1751
12.117 | -34.564 | -20.1839
12.1231 | -34.5645 | -20.1928
12.1292 | -34.565 | -20.2015
12.1353 | -34.5654 | -20.2106
12.1414 | -34.5659 | -20.2193
12.1475 | -34.5664 | -20.2281
12.1536 | -34.5668 | -20.2371
12.1597 | -34.5673 | -20.2458
12.1658 | -34.5677 | -20.2547
12.1719 | -34.5682 | -20.2633
12.178 | -34.5686 | -20.2721
12.1841 | -34.5691 | -20.2809
12.1902 | -34.5696 | -20.2899
12.1963 | -34.5701 | -20.2987
12.2024 | -34.5705 | -20.3074
12.2085 | -34.571 | -20.3163
12.2146 | -34.5714 | -20.3252
12.2207 | -34.5719 | -20.334
12.2268 | -34.5724 | -20.3427
12.2329 | -34.5728 | -20.3514
12.239 | -34.5733 | -20.3601
12.2451 | -34.5738 | -20.3689
12.2513 | -34.5742 | -20.3776
12.2574 | -34.5746 | -20.3865
12.2635 | -34.5751 | -20.3953
12.2696 | -34.5756 | -20.404
12.2757 | -34.576 | -20.4127
12.2818 | -34.5765 | -20.4216
12.2879 | -34.577 | -20.4305
12.294 | -34.5775 | -20.4392
12.3001 | -34.5779 | -20.448
12.3062 | -34.5784 | -20.4569
12.3123 | -34.5789 | -20.4654
12.3184 | -34.5793 | -20.474
12.3245 | -34.5798 | -20.4827
12.3306 | -34.5803 | -20.4915
12.3367 | -34.5807 | -20.5003
12.3428 | -34.5812 | -20.5091
12.3489 | -34.5817 | -20.5177
12.355 | -34.5821 | -20.5266
12.3611 | -34.5826 | -20.5354
12.3672 | -34.583 | -20.5441
12.3733 | -34.5835 | -20.5529
12.3794 | -34.584 | -20.5616
12.3855 | -34.5844 | -20.5703
12.3916 | -34.5849 | -20.579
12.3978 | -34.5854 | -20.588
12.4039 | -34.5858 | -20.597
12.41 | -34.5863 | -20.6056
12.4161 | -34.5867 | -20.6142
12.4222 | -34.5872 | -20.6229
12.4283 | -34.5877 | -20.6316
12.4344 | -34.5881 | -20.6406
12.4405 | -34.5886 | -20.6494
12.4466 | -34.589 | -20.6581
12.4527 | -34.5895 | -20.6668
12.4588 | -34.5899 | -20.6756
12.4649 | -34.5904 | -20.6844
12.471 | -34.5909 | -20.6934
12.4771 | -34.5913 | -20.702
12.4832 | -34.5918 | -20.7107
12.4893 | -34.5922 | -20.7193
12.4954 | -34.5927 | -20.7278
12.5015 | -34.5931 | -20.7366
12.5076 | -34.5936 | -20.7453
12.5137 | -34.5941 | -20.7541
12.5198 | -34.5945 | -20.7627
12.5259 | -34.5949 | -20.7714
12.532 | -34.5954 | -20.7803
12.5382 | -34.5958 | -20.7891
12.5443 | -34.5963 | -20.7979
12.5504 | -34.5967 | -20.8066
12.5565 | -34.5972 | -20.8154
12.5626 | -34.5977 | -20.8244
12.5687 | -34.5982 | -20.8332
12.5748 | -34.5986 | -20.842
12.5809 | -34.5991 | -20.8506
12.587 | -34.5995 | -20.8596
12.5931 | -34.6 | -20.8682
12.5992 | -34.6005 | -20.877
12.6053 | -34.6009 | -20.8856
12.6114 | -34.6014 | -20.8943
12.6175 | -34.6019 | -20.9033
12.6236 | -34.6023 | -20.9121
12.6297 | -34.6028 | -20.9208
12.6358 | -34.6032 | -20.9295
12.6419 | -34.6037 | -20.9382
12.648 | -34.6041 | -20.9472
12.6541 | -34.6046 | -20.9558
12.6602 | -34.605 | -20.9642
12.6663 | -34.6055 | -20.9729
12.6724 | -34.606 | -20.9815
12.6785 | -34.6064 | -20.9903
12.6847 | -34.6069 | -20.9989
12.6908 | -34.6073 | -21.0076
12.6969 | -34.6078 | -21.0164
12.703 | -34.6082 | -21.0252
12.7091 | -34.6086 | -21.0338
12.7152 | -34.609 | -21.0424
12.7213 | -34.6095 | -21.0511
12.7274 | -34.6099 | -21.0599
12.7335 | -34.6104 | -21.0682
12.7396 | -34.6109 | -21.0767
12.7457 | -34.6113 | -21.0855
12.7518 | -34.6118 | -21.0942
12.7579 | -34.6122 | -21.1031
12.764 | -34.6127 | -21.1118
12.7701 | -34.6131 | -21.1207
12.7762 | -34.6136 | -21.1293
12.7823 | -34.614 | -21.1381
12.7884 | -34.6145 | -21.147
12.7945 | -34.6149 | -21.1558
12.8006 | -34.6154 | -21.1645
12.8067 | -34.6159 | -21.173
12.8128 | -34.6163 | -21.1816
12.8189 | -34.6168 | -21.1902
12.8251 | -34.6172 | -21.1989
12.8312 | -34.6177 | -21.2078
12.8373 | -34.6181 | -21.2164
12.8434 | -34.6186 | -21.2251
12.8495 | -34.619 | -21.2337
12.8556 | -34.6195 | -21.2422
12.8617 | -34.6199 | -21.2511
12.8678 | -34.6204 | -21.2599
12.8739 | -34.6208 | -21.2685
12.88 | -34.6212 | -21.2769
12.8861 | -34.6217 | -21.2855
12.8922 | -34.6221 | -21.2942
12.8983 | -34.6226 | -21.303
12.9044 | -34.623 | -21.3115
12.9105 | -34.6235 | -21.3201
12.9166 | -34.6239 | -21.3289
12.9227 | -34.6244 | -21.3376
12.9288 | -34.6248 | -21.3462
12.9349 | -34.6253 | -21.3549
12.941 | -34.6258 | -21.3636
12.9471 | -34.6262 | -21.3722
12.9532 | -34.6267 | -21.3809
12.9593 | -34.6271 | -21.3897
12.9654 | -34.6276 | -21.3984
12.9716 | -34.628 | -21.4072
12.9777 | -34.6285 | -21.416
12.9838 | -34.629 | -21.4246
12.9899 | -34.6294 | -21.4333
12.996 | -34.6298 | -21.4418
13.0021 | -34.6303 | -21.4505
13.0082 | -34.6308 | -21.459
13.0143 | -34.6313 | -21.4673
13.0204 | -34.6317 | -21.4759
13.0265 | -34.6322 | -21.4846
13.0326 | -34.6326 | -21.4932
13.0387 | -34.633 | -21.5019
13.0448 | -34.6335 | -21.5106
13.0509 | -34.6339 | -21.5192
13.057 | -34.6344 | -21.5278
13.0631 | -34.6349 | -21.5363
13.0692 | -34.6354 | -21.5448
13.0753 | -34.6358 | -21.5536
13.0814 | -34.6363 | -21.5624
13.0875 | -34.6367 | -21.5708
13.0936 | -34.6371 | -21.5795
13.0997 | -34.6376 | -21.588
13.1058 | -34.6381 | -21.5964
13.112 | -34.6385 | -21.6051
13.1181 | -34.639 | -21.6135
13.1242 | -34.6394 | -21.6221
13.1303 | -34.6399 | -21.6307
13.1364 | -34.6403 | -21.6392
13.1425 | -34.6407 | -21.6477
13.1486 | -34.6412 | -21.6562
13.1547 | -34.6416 | -21.6649
13.1608 | -34.6421 | -21.6735
13.1669 | -34.6426 | -21.6821
13.173 | -34.643 | -21.6909
13.1791 | -34.6435 | -21.6993
13.1852 | -34.6439 | -21.708
13.1913 | -34.6444 | -21.7165
13.1974 | -34.6448 | -21.7253
13.2035 | -34.6452 | -21.7339
13.2096 | -34.6457 | -21.7424
13.2157 | -34.6462 | -21.751
13.2218 | -34.6466 | -21.7594
13.2279 | -34.6471 | -21.768
13.234 | -34.6476 | -21.7766
13.2401 | -34.648 | -21.7851
13.2462 | -34.6484 | -21.7937
13.2524 | -34.6489 | -21.8024
13.2585 | -34.6493 | -21.811
13.2646 | -34.6498 | -21.8196
13.2707 | -34.6503 | -21.8283
13.2768 | -34.6507 | -21.8371
13.2829 | -34.6512 | -21.8459
13.289 | -34.6516 | -21.8544
13.2951 | -34.6521 | -21.8628
13.3012 | -34.6525 | -21.8713
13.3073 | -34.653 | -21.88
13.3134 | -34.6534 | -21.8886
13.3195 | -34.6538 | -21.897
13.3256 | -34.6542 | -21.9055
13.3317 | -34.6547 | -21.9139
13.3378 | -34.6551 | -21.9225
13.3439 | -34.6556 | -21.9312
13.35 | -34.656 | -21.9399
13.3561 | -34.6565 | -21.9482
13.3622 | -34.657 | -21.9565
13.3683 | -34.6574 | -21.9651
13.3744 | -34.6578 | -21.9735
13.3805 | -34.6583 | -21.982
13.3866 | -34.6587 | -21.9904
13.3927 | -34.6592 | -21.999
13.3989 | -34.6596 | -22.0076
13.405 | -34.6601 | -22.0162
13.4111 | -34.6605 | -22.0246
13.4172 | -34.6609 | -22.033
13.4233 | -34.6614 | -22.0418
13.4294 | -34.6618 | -22.0505
13.4355 | -34.6623 | -22.0589
13.4416 | -34.6627 | -22.0673
13.4477 | -34.6632 | -22.0759
13.4538 | -34.6637 | -22.0845
13.4599 | -34.6641 | -22.0931
13.466 | -34.6646 | -22.1018
13.4721 | -34.665 | -22.1102
13.4782 | -34.6655 | -22.1188
13.4843 | -34.6659 | -22.1276
13.4904 | -34.6664 | -22.136
13.4965 | -34.6668 | -22.1447
13.5026 | -34.6672 | -22.1534
13.5087 | -34.6677 | -22.1618
13.5148 | -34.6681 | -22.1705
13.5209 | -34.6685 | -22.1791
13.527 | -34.669 | -22.1877
13.5331 | -34.6695 | -22.1962
13.5393 | -34.6699 | -22.2046
13.5454 | -34.6704 | -22.2131
13.5515 | -34.6708 | -22.2218
13.5576 | -34.6713 | -22.2306
13.5637 | -34.6717 | -22.2393
13.5698 | -34.6722 | -22.2477
13.5759 | -34.6726 | -22.2558
13.582 | -34.6731 | -22.2645
13.5881 | -34.6735 | -22.2728
13.5942 | -34.674 | -22.2816
13.6003 | -34.6745 | -22.2901
13.6064 | -34.6749 | -22.2987
13.6125 | -34.6754 | -22.3072
13.6186 | -34.6758 | -22.3157
13.6247 | -34.6763 | -22.3244
13.6308 | -34.6768 | -22.333
13.6369 | -34.6772 | -22.3417
13.643 | -34.6776 | -22.3504
13.6491 | -34.6781 | -22.3589
13.6552 | -34.6786 | -22.3675
13.6613 | -34.679 | -22.3759
13.6674 | -34.6794 | -22.3844
13.6735 | -34.6799 | -22.393
13.6796 | -34.6804 | -22.4015
13.6858 | -34.6808 | -22.4098
13.6919 | -34.6813 | -22.4183
13.698 | -34.6817 | -22.4269
13.7041 | -34.6822 | -22.4357
13.7102 | -34.6826 | -22.4442
13.7163 | -34.6831 | -22.4525
13.7224 | -34.6835 | -22.4609
13.7285 | -34.684 | -22.4696
13.7346 | -34.6844 | -22.478
13.7407 | -34.6849 | -22.4865
13.7468 | -34.6853 | -22.495
13.7529 | -34.6858 | -22.5033
13.759 | -34.6862 | -22.512
13.7651 | -34.6867 | -22.5205
13.7712 | -34.6871 | -22.5291
13.7773 | -34.6876 | -22.5377
13.7834 | -34.688 | -22.5461
13.7895 | -34.6885 | -22.5546
13.7956 | -34.6889 | -22.5632
13.8017 | -34.6894 | -22.5715
13.8078 | -34.6898 | -22.58
13.8139 | -34.6903 | -22.5883
13.82 | -34.6907 | -22.5967
13.8262 | -34.6912 | -22.6052
13.8323 | -34.6917 | -22.6139
13.8384 | -34.6921 | -22.6225
13.8445 | -34.6926 | -22.6309
13.8506 | -34.693 | -22.6393
13.8567 | -34.6935 | -22.6477
13.8628 | -34.694 | -22.6564
13.8689 | -34.6944 | -22.665
13.875 | -34.6949 | -22.6735
13.8811 | -34.6953 | -22.682
13.8872 | -34.6957 | -22.6905
13.8933 | -34.6961 | -22.6988
13.8994 | -34.6966 | -22.7072
13.9055 | -34.6971 | -22.7158
13.9116 | -34.6975 | -22.7241
13.9177 | -34.698 | -22.7326
13.9238 | -34.6984 | -22.7408
13.9299 | -34.6989 | -22.7492
13.936 | -34.6993 | -22.7578
13.9421 | -34.6997 | -22.7661
13.9482 | -34.7002 | -22.7744
13.9543 | -34.7006 | -22.783
13.9604 | -34.7011 | -22.7915
13.9665 | -34.7015 | -22.7998
13.9727 | -34.702 | -22.8085
13.9788 | -34.7025 | -22.8169
13.9849 | -34.7029 | -22.8255
13.991 | -34.7034 | -22.8342
13.9971 | -34.7038 | -22.8425
14.0032 | -34.7043 | -22.8511
14.0093 | -34.7047 | -22.8595
14.0154 | -34.7052 | -22.8679
14.0215 | -34.7056 | -22.8761
14.0276 | -34.7061 | -22.8845
14.0337 | -34.7066 | -22.8929
14.0398 | -34.7071 | -22.9013
14.0459 | -34.7076 | -22.9098
14.052 | -34.708 | -22.9185
14.0581 | -34.7084 | -22.9271
14.0642 | -34.7089 | -22.9356
14.0703 | -34.7094 | -22.944
14.0764 | -34.7098 | -22.9524
14.0825 | -34.7103 | -22.9608
14.0886 | -34.7107 | -22.9693
14.0947 | -34.7112 | -22.9777
14.1008 | -34.7116 | -22.986
14.1069 | -34.7121 | -22.9942
14.1131 | -34.7125 | -23.0028
14.1192 | -34.713 | -23.0111
14.1253 | -34.7134 | -23.0194
14.1314 | -34.7139 | -23.0278
14.1375 | -34.7144 | -23.0361
14.1436 | -34.7149 | -23.0444
14.1497 | -34.7153 | -23.0527
14.1558 | -34.7158 | -23.0613
14.1619 | -34.7162 | -23.0697
14.168 | -34.7167 | -23.0781
14.1741 | -34.7172 | -23.0865
14.1802 | -34.7176 | -23.0949
14.1863 | -34.7181 | -23.1034
14.1924 | -34.7186 | -23.1118
14.1985 | -34.7191 | -23.12
14.2046 | -34.7195 | -23.1286
14.2107 | -34.72 | -23.1367
14.2168 | -34.7204 | -23.1451
14.2229 | -34.7209 | -23.1536
14.229 | -34.7213 | -23.162
14.2351 | -34.7218 | -23.1704
14.2412 | -34.7222 | -23.1788
14.2473 | -34.7227 | -23.1871
14.2534 | -34.7232 | -23.1954
14.2596 | -34.7236 | -23.2037
14.2657 | -34.724 | -23.2121
14.2718 | -34.7245 | -23.2205
14.2779 | -34.725 | -23.2289
14.284 | -34.7255 | -23.2373
14.2901 | -34.7259 | -23.2457
14.2962 | -34.7264 | -23.2541
14.3023 | -34.7269 | -23.2625
14.3084 | -34.7273 | -23.2708
14.3145 | -34.7278 | -23.2792
14.3206 | -34.7283 | -23.288
14.3267 | -34.7288 | -23.2965
14.3328 | -34.7292 | -23.3048
14.3389 | -34.7297 | -23.3132
14.345 | -34.7301 | -23.3216
14.3511 | -34.7306 | -23.3299
14.3572 | -34.7311 | -23.3384
14.3633 | -34.7315 | -23.3469
14.3694 | -34.732 | -23.3552
14.3755 | -34.7325 | -23.3635
14.3816 | -34.733 | -23.3721
14.3877 | -34.7335 | -23.3802
14.3938 | -34.734 | -23.3888
14.4 | -34.7344 | -23.3971
14.4061 | -34.7349 | -23.4055
14.4122 | -34.7353 | -23.4139
14.4183 | -34.7358 | -23.4222
14.4244 | -34.7363 | -23.4305
14.4305 | -34.7368 | -23.4388
14.4366 | -34.7373 | -23.4471
14.4427 | -34.7377 | -23.4554
14.4488 | -34.7382 | -23.464
14.4549 | -34.7387 | -23.4723
14.461 | -34.7391 | -23.4805
14.4671 | -34.7396 | -23.4889
14.4732 | -34.74 | -23.4975
14.4793 | -34.7405 | -23.5058
14.4854 | -34.741 | -23.5141
14.4915 | -34.7415 | -23.5222
14.4976 | -34.7419 | -23.5303
14.5037 | -34.7424 | -23.5385
14.5098 | -34.7429 | -23.547
14.5159 | -34.7434 | -23.5555
14.522 | -34.7438 | -23.5639
14.5281 | -34.7443 | -23.5721
14.5342 | -34.7448 | -23.5805
14.5403 | -34.7452 | -23.5888
14.5465 | -34.7457 | -23.5971
14.5526 | -34.7462 | -23.6057
14.5587 | -34.7467 | -23.6142
14.5648 | -34.7472 | -23.6224
14.5709 | -34.7476 | -23.6305
14.577 | -34.7481 | -23.639
14.5831 | -34.7485 | -23.6474
14.5892 | -34.749 | -23.6558
14.5953 | -34.7495 | -23.6638
14.6014 | -34.7499 | -23.6722
14.6075 | -34.7504 | -23.6804
14.6136 | -34.7508 | -23.6887
14.6197 | -34.7513 | -23.6971
14.6258 | -34.7517 | -23.7054
14.6319 | -34.7522 | -23.7135
14.638 | -34.7526 | -23.7218
14.6441 | -34.7531 | -23.7303
14.6502 | -34.7536 | -23.7388
14.6563 | -34.754 | -23.7471
14.6624 | -34.7545 | -23.7555
14.6685 | -34.755 | -23.7636
14.6746 | -34.7555 | -23.7719
14.6807 | -34.7559 | -23.7802
14.6869 | -34.7564 | -23.7882
14.693 | -34.7569 | -23.7964
14.6991 | -34.7574 | -23.8049
14.7052 | -34.7579 | -23.813
14.7113 | -34.7584 | -23.8213
14.7174 | -34.7588 | -23.8296
14.7235 | -34.7593 | -23.8378
14.7296 | -34.7597 | -23.8462
14.7357 | -34.7603 | -23.8544
14.7418 | -34.7607 | -23.8625
14.7479 | -34.7612 | -23.8708
14.754 | -34.7616 | -23.8793
14.7601 | -34.7621 | -23.8878
14.7662 | -34.7625 | -23.8962
14.7723 | -34.763 | -23.9042
14.7784 | -34.7635 | -23.9126
14.7845 | -34.7639 | -23.9208
14.7906 | -34.7644 | -23.9291
14.7967 | -34.7649 | -23.9372
14.8028 | -34.7653 | -23.9454
14.8089 | -34.7658 | -23.9535
14.815 | -34.7663 | -23.9617
14.8211 | -34.7668 | -23.9699
14.8272 | -34.7673 | -23.9782
14.8334 | -34.7677 | -23.9863
14.8395 | -34.7682 | -23.9946
14.8456 | -34.7687 | -24.0027
14.8517 | -34.7692 | -24.0111
14.8578 | -34.7697 | -24.0194
14.8639 | -34.7702 | -24.0278
14.87 | -34.7707 | -24.0359
14.8761 | -34.7711 | -24.0439
14.8822 | -34.7717 | -24.0524
14.8883 | -34.7721 | -24.0608
14.8944 | -34.7726 | -24.0692
14.9005 | -34.7731 | -24.0773
14.9066 | -34.7736 | -24.0856
14.9127 | -34.7741 | -24.0936
14.9188 | -34.7745 | -24.1019
14.9249 | -34.775 | -24.1102
14.931 | -34.7755 | -24.1181
14.9371 | -34.776 | -24.1264
14.9432 | -34.7765 | -24.1348
14.9493 | -34.777 | -24.1429
14.9554 | -34.7775 | -24.1512
14.9615 | -34.7779 | -24.1596
14.9676 | -34.7784 | -24.1677
14.9738 | -34.7789 | -24.176
14.9799 | -34.7793 | -24.1842
14.986 | -34.7798 | -24.1922
14.9921 | -34.7803 | -24.2004
14.9982 | -34.7808 | -24.2087
15.0043 | -34.7812 | -24.217
15.0104 | -34.7817 | -24.225
15.0165 | -34.7822 | -24.2332
15.0226 | -34.7827 | -24.2411
15.0287 | -34.7832 | -24.2491
15.0348 | -34.7836 | -24.2574
15.0409 | -34.7841 | -24.2656
15.047 | -34.7846 | -24.274
15.0531 | -34.7851 | -24.282
15.0592 | -34.7856 | -24.2904
15.0653 | -34.786 | -24.2985
15.0714 | -34.7865 | -24.3068
15.0775 | -34.787 | -24.315
15.0836 | -34.7875 | -24.3233
15.0897 | -34.788 | -24.3312
15.0958 | -34.7885 | -24.3392
15.1019 | -34.789 | -24.3474
15.108 | -34.7895 | -24.3554
15.1141 | -34.79 | -24.3638
15.1203 | -34.7905 | -24.3717
15.1264 | -34.791 | -24.3798
15.1325 | -34.7914 | -24.3878
15.1386 | -34.7919 | -24.396
15.1447 | -34.7924 | -24.4044
15.1508 | -34.7929 | -24.4127
15.1569 | -34.7934 | -24.4207
15.163 | -34.7939 | -24.4288
15.1691 | -34.7944 | -24.437
15.1752 | -34.7948 | -24.4453
15.1813 | -34.7953 | -24.4534
15.1874 | -34.7959 | -24.4616
15.1935 | -34.7964 | -24.4695
15.1996 | -34.7968 | -24.4776
15.2057 | -34.7973 | -24.4858
15.2118 | -34.7978 | -24.4939
15.2179 | -34.7983 | -24.5022
15.224 | -34.7988 | -24.5103
15.2301 | -34.7993 | -24.5181
15.2362 | -34.7998 | -24.5263
15.2423 | -34.8003 | -24.5346
15.2484 | -34.8008 | -24.5427
15.2545 | -34.8013 | -24.5508
15.2607 | -34.8018 | -24.5591
15.2668 | -34.8023 | -24.5671
15.2729 | -34.8028 | -24.5755
15.279 | -34.8033 | -24.5838
15.2851 | -34.8038 | -24.5921
15.2912 | -34.8043 | -24.6001
15.2973 | -34.8048 | -24.6083
15.3034 | -34.8053 | -24.6163
15.3095 | -34.8058 | -24.6245
15.3156 | -34.8062 | -24.6327
15.3217 | -34.8067 | -24.6409
15.3278 | -34.8072 | -24.649
15.3339 | -34.8077 | -24.657
15.34 | -34.8082 | -24.665
15.3461 | -34.8087 | -24.6732
15.3522 | -34.8092 | -24.6812
15.3583 | -34.8096 | -24.6897
15.3644 | -34.8101 | -24.6973
15.3705 | -34.8107 | -24.7051
15.3766 | -34.8112 | -24.7133
15.3827 | -34.8116 | -24.7214
15.3888 | -34.8121 | -24.7294
15.3949 | -34.8126 | -24.7374
15.401 | -34.813 | -24.7453
15.4072 | -34.8135 | -24.7531
15.4133 | -34.814 | -24.7611
15.4194 | -34.8145 | -24.7693
15.4255 | -34.8149 | -24.7774
15.4316 | -34.8154 | -24.7854
15.4377 | -34.8159 | -24.7934
15.4438 | -34.8164 | -24.8013
15.4499 | -34.8169 | -24.8094
15.456 | -34.8174 | -24.8175
15.4621 | -34.8179 | -24.8253
15.4682 | -34.8184 | -24.8335
15.4743 | -34.8189 | -24.8413
15.4804 | -34.8193 | -24.8494
15.4865 | -34.8198 | -24.8575
15.4926 | -34.8203 | -24.8654
15.4987 | -34.8208 | -24.8736
15.5048 | -34.8213 | -24.8814
15.5109 | -34.8218 | -24.8893
15.517 | -34.8222 | -24.8975
15.5231 | -34.8228 | -24.9054
15.5292 | -34.8233 | -24.913
15.5353 | -34.8237 | -24.9211
15.5414 | -34.8242 | -24.929
15.5476 | -34.8247 | -24.9368
15.5537 | -34.8252 | -24.945
15.5598 | -34.8257 | -24.953
15.5659 | -34.8262 | -24.961
15.572 | -34.8267 | -24.9692
15.5781 | -34.8272 | -24.9772
15.5842 | -34.8277 | -24.9851
15.5903 | -34.8283 | -24.9931
15.5964 | -34.8288 | -25.0008
15.6025 | -34.8293 | -25.0089
15.6086 | -34.8298 | -25.0168
15.6147 | -34.8303 | -25.0249
15.6208 | -34.8308 | -25.0327
15.6269 | -34.8313 | -25.0406
15.633 | -34.8318 | -25.0485
15.6391 | -34.8323 | -25.0567
15.6452 | -34.8328 | -25.0648
15.6513 | -34.8333 | -25.0727
15.6574 | -34.8338 | -25.0812
15.6635 | -34.8343 | -25.0892
15.6696 | -34.8348 | -25.097
15.6757 | -34.8353 | -25.1048
15.6818 | -34.8358 | -25.1124
15.688 | -34.8363 | -25.1203
15.6941 | -34.8368 | -25.1286
15.7002 | -34.8372 | -25.1362
15.7063 | -34.8378 | -25.144
15.7124 | -34.8382 | -25.1522
15.7185 | -34.8387 | -25.1598
15.7246 | -34.8392 | -25.1679
15.7307 | -34.8398 | -25.1756
15.7368 | -34.8403 | -25.1834
15.7429 | -34.8408 | -25.1915
15.749 | -34.8413 | -25.1994
15.7551 | -34.8418 | -25.2071
15.7612 | -34.8424 | -25.215
15.7673 | -34.8429 | -25.2226
15.7734 | -34.8433 | -25.2303
15.7795 | -34.8438 | -25.2383
15.7856 | -34.8443 | -25.2459
15.7917 | -34.8448 | -25.2537
15.7978 | -34.8453 | -25.2618
15.8039 | -34.8459 | -25.2696
15.81 | -34.8464 | -25.2779
15.8161 | -34.8469 | -25.2859
15.8222 | -34.8474 | -25.2937
15.8283 | -34.8479 | -25.3015
15.8345 | -34.8484 | -25.3094
15.8406 | -34.8489 | -25.3173
15.8467 | -34.8495 | -25.3252
15.8528 | -34.85 | -25.3331
15.8589 | -34.8505 | -25.341
15.865 | -34.851 | -25.3487
15.8711 | -34.8515 | -25.3566
15.8772 | -34.852 | -25.3645
15.8833 | -34.8524 | -25.3727
15.8894 | -34.853 | -25.3806
15.8955 | -34.8534 | -25.3885
15.9016 | -34.8539 | -25.3962
15.9077 | -34.8545 | -25.4043
15.9138 | -34.8549 | -25.4125
15.9199 | -34.8554 | -25.4203
15.926 | -34.8559 | -25.4281
15.9321 | -34.8564 | -25.4359
15.9382 | -34.8569 | -25.4439
15.9443 | -34.8574 | -25.4519
15.9504 | -34.8579 | -25.4596
15.9565 | -34.8584 | -25.4673
15.9626 | -34.8588 | -25.4753
15.9687 | -34.8593 | -25.4831
15.9749 | -34.8598 | -25.4907
15.981 | -34.8603 | -25.4988
15.9871 | -34.8608 | -25.5068
15.9932 | -34.8613 | -25.5146
15.9993 | -34.8618 | -25.5228
16.0054 | -34.8623 | -25.5306
16.0115 | -34.8628 | -25.5388
16.0176 | -34.8632 | -25.5468
16.0237 | -34.8638 | -25.5549
16.0298 | -34.8643 | -25.5627
16.0359 | -34.8648 | -25.5706
16.042 | -34.8653 | -25.5785
16.0481 | -34.8658 | -25.5864
16.0542 | -34.8663 | -25.5944
16.0603 | -34.8669 | -25.6021
16.0664 | -34.8674 | -25.6099
16.0725 | -34.8679 | -25.6175
16.0786 | -34.8684 | -25.625
16.0847 | -34.8689 | -25.6333
16.0908 | -34.8694 | -25.6415
16.0969 | -34.8699 | -25.6496
16.103 | -34.8704 | -25.6578
16.1091 | -34.8709 | -25.6655
16.1152 | -34.8714 | -25.6736
16.1214 | -34.8719 | -25.6816
16.1275 | -34.8724 | -25.6894
16.1336 | -34.8729 | -25.6971
16.1397 | -34.8735 | -25.7047
16.1458 | -34.874 | -25.7123
16.1519 | -34.8745 | -25.7204
16.158 | -34.875 | -25.7283
16.1641 | -34.8755 | -25.7361
16.1702 | -34.8759 | -25.744
16.1763 | -34.8764 | -25.752
16.1824 | -34.8769 | -25.7594
16.1885 | -34.8774 | -25.767
16.1946 | -34.8779 | -25.7748
16.2007 | -34.8784 | -25.7828
16.2068 | -34.8789 | -25.7908
16.2129 | -34.8794 | -25.7987
16.219 | -34.8799 | -25.8064
16.2251 | -34.8805 | -25.8139
16.2312 | -34.881 | -25.8218
16.2373 | -34.8815 | -25.8297
16.2434 | -34.882 | -25.8373
16.2495 | -34.8825 | -25.845
16.2556 | -34.883 | -25.8533
16.2618 | -34.8836 | -25.8611
16.2679 | -34.8841 | -25.8692
16.274 | -34.8846 | -25.8769
16.2801 | -34.885 | -25.8845
16.2862 | -34.8856 | -25.8922
16.2923 | -34.8861 | -25.8999
16.2984 | -34.8866 | -25.9077
16.3045 | -34.8871 | -25.9156
16.3106 | -34.8876 | -25.9233
16.3167 | -34.8881 | -25.9313
16.3228 | -34.8886 | -25.9389
16.3289 | -34.8891 | -25.9469
16.335 | -34.8896 | -25.9544
16.3411 | -34.8901 | -25.9624
16.3472 | -34.8906 | -25.9702
16.3533 | -34.8911 | -25.9782
16.3594 | -34.8916 | -25.9856
16.3655 | -34.8921 | -25.9931
16.3716 | -34.8926 | -26.001
16.3777 | -34.8932 | -26.0088
16.3838 | -34.8937 | -26.0165
16.3899 | -34.8942 | -26.0242
16.396 | -34.8947 | -26.0319
16.4021 | -34.8952 | -26.0394
16.4083 | -34.8957 | -26.0471
16.4144 | -34.8962 | -26.055
16.4205 | -34.8967 | -26.0628
16.4266 | -34.8972 | -26.0704
16.4327 | -34.8977 | -26.078
16.4388 | -34.8982 | -26.0859
16.4449 | -34.8987 | -26.0936
16.451 | -34.8992 | -26.1014
16.4571 | -34.8998 | -26.1093
16.4632 | -34.9003 | -26.1168
16.4693 | -34.9007 | -26.1246
16.4754 | -34.9012 | -26.1324
16.4815 | -34.9017 | -26.1401
16.4876 | -34.9022 | -26.1477
16.4937 | -34.9027 | -26.1554
16.4998 | -34.9033 | -26.1629
16.5059 | -34.9038 | -26.1707
16.512 | -34.9043 | -26.1788
16.5181 | -34.9049 | -26.1869
16.5242 | -34.9054 | -26.1949
16.5303 | -34.9059 | -26.2024
16.5364 | -34.9064 | -26.2099
16.5425 | -34.9069 | -26.2175
16.5487 | -34.9074 | -26.2252
16.5548 | -34.9079 | -26.2331
16.5609 | -34.9084 | -26.2409
16.567 | -34.909 | -26.2484
16.5731 | -34.9095 | -26.2561
16.5792 | -34.91 | -26.2637
16.5853 | -34.9105 | -26.2716
16.5914 | -34.9109 | -26.2795
16.5975 | -34.9114 | -26.2873
16.6036 | -34.912 | -26.2949
16.6097 | -34.9124 | -26.3025
16.6158 | -34.913 | -26.3102
16.6219 | -34.9135 | -26.3181
16.628 | -34.914 | -26.3258
16.6341 | -34.9145 | -26.3333
16.6402 | -34.915 | -26.341
16.6463 | -34.9154 | -26.3487
16.6524 | -34.9159 | -26.3565
16.6585 | -34.9165 | -26.3641
16.6646 | -34.917 | -26.3719
16.6707 | -34.9175 | -26.3794
16.6768 | -34.918 | -26.3871
16.6829 | -34.9185 | -26.3947
16.689 | -34.919 | -26.4025
16.6952 | -34.9195 | -26.4101
16.7013 | -34.92 | -26.418
16.7074 | -34.9205 | -26.4254
16.7135 | -34.921 | -26.4331
16.7196 | -34.9215 | -26.441
16.7257 | -34.922 | -26.4486
16.7318 | -34.9225 | -26.4562
16.7379 | -34.923 | -26.4636
16.744 | -34.9235 | -26.4712
16.7501 | -34.924 | -26.4789
16.7562 | -34.9245 | -26.4866
16.7623 | -34.925 | -26.4941
16.7684 | -34.9255 | -26.502
16.7745 | -34.926 | -26.5097
16.7806 | -34.9266 | -26.517
16.7867 | -34.9271 | -26.5246
16.7928 | -34.9276 | -26.5321
16.7989 | -34.9281 | -26.5398
16.805 | -34.9286 | -26.5474
16.8111 | -34.9291 | -26.555
16.8172 | -34.9296 | -26.5627
16.8233 | -34.9301 | -26.5705
16.8294 | -34.9306 | -26.5782
16.8356 | -34.9311 | -26.5859
16.8417 | -34.9316 | -26.5935
16.8478 | -34.9321 | -26.6011
16.8539 | -34.9325 | -26.6087
16.86 | -34.933 | -26.6165
16.8661 | -34.9336 | -26.6242
16.8722 | -34.9341 | -26.6319
16.8783 | -34.9346 | -26.6396
16.8844 | -34.9351 | -26.6472
16.8905 | -34.9356 | -26.6548
16.8966 | -34.9361 | -26.6623
16.9027 | -34.9367 | -26.67
16.9088 | -34.9371 | -26.6777
16.9149 | -34.9376 | -26.6853
16.921 | -34.9382 | -26.6929
16.9271 | -34.9387 | -26.7004
16.9332 | -34.9392 | -26.708
16.9393 | -34.9397 | -26.7157
16.9454 | -34.9402 | -26.7231
16.9515 | -34.9407 | -26.7307
16.9576 | -34.9412 | -26.7381
16.9637 | -34.9417 | -26.7454
16.9698 | -34.9422 | -26.7529
16.9759 | -34.9427 | -26.7606
16.9821 | -34.9433 | -26.7684
16.9882 | -34.9438 | -26.7757
16.9943 | -34.9442 | -26.7832
17.0004 | -34.9448 | -26.7907
17.0065 | -34.9453 | -26.7985
17.0126 | -34.9458 | -26.8061
17.0187 | -34.9463 | -26.8137
17.0248 | -34.9468 | -26.821
17.0309 | -34.9473 | -26.8288
17.037 | -34.9478 | -26.8364
17.0431 | -34.9483 | -26.8439
17.0492 | -34.9488 | -26.8514
17.0553 | -34.9493 | -26.8587
17.0614 | -34.9498 | -26.8662
17.0675 | -34.9503 | -26.8736
17.0736 | -34.9508 | -26.8814
17.0797 | -34.9513 | -26.8888
17.0858 | -34.9518 | -26.8964
17.0919 | -34.9523 | -26.9039
17.098 | -34.9529 | -26.9113
17.1041 | -34.9534 | -26.9188
17.1102 | -34.9539 | -26.9264
17.1163 | -34.9544 | -26.9341
17.1225 | -34.9549 | -26.9413
17.1286 | -34.9554 | -26.9485
17.1347 | -34.9559 | -26.9561
17.1408 | -34.9564 | -26.9637
17.1469 | -34.9569 | -26.9713
17.153 | -34.9574 | -26.9789
17.1591 | -34.9579 | -26.9865
17.1652 | -34.9584 | -26.9939
17.1713 | -34.959 | -27.0012
17.1774 | -34.9595 | -27.0088
17.1835 | -34.96 | -27.0162
17.1896 | -34.9605 | -27.0235
17.1957 | -34.961 | -27.0311
17.2018 | -34.9615 | -27.0383
17.2079 | -34.962 | -27.0458
17.214 | -34.9625 | -27.0535
17.2201 | -34.963 | -27.061
17.2262 | -34.9635 | -27.0686
17.2323 | -34.964 | -27.0762
17.2384 | -34.9646 | -27.0835
17.2445 | -34.9651 | -27.0908
17.2506 | -34.9656 | -27.0983
17.2567 | -34.9661 | -27.1058
17.2628 | -34.9666 | -27.1133
17.269 | -34.9671 | -27.1206
17.2751 | -34.9676 | -27.1281
17.2812 | -34.9681 | -27.1357
17.2873 | -34.9686 | -27.1432
17.2934 | -34.9691 | -27.1508
17.2995 | -34.9696 | -27.1584
17.3056 | -34.9701 | -27.1656
17.3117 | -34.9706 | -27.1732
17.3178 | -34.9711 | -27.1806
17.3239 | -34.9716 | -27.1881
17.33 | -34.9721 | -27.1956
17.3361 | -34.9726 | -27.203
17.3422 | -34.9731 | -27.2101
17.3483 | -34.9736 | -27.2175
17.3544 | -34.9741 | -27.225
17.3605 | -34.9746 | -27.2324
17.3666 | -34.9751 | -27.2398
17.3727 | -34.9756 | -27.2472
17.3788 | -34.9761 | -27.2542
17.3849 | -34.9766 | -27.2616
17.391 | -34.9771 | -27.2691
17.3971 | -34.9776 | -27.2765
17.4032 | -34.9781 | -27.2841
17.4094 | -34.9786 | -27.2913
17.4155 | -34.9791 | -27.2987
17.4216 | -34.9796 | -27.3061
17.4277 | -34.9801 | -27.3137
17.4338 | -34.9806 | -27.3212
17.4399 | -34.9812 | -27.3287
17.446 | -34.9817 | -27.3359
17.4521 | -34.9821 | -27.3434
17.4582 | -34.9826 | -27.3507
17.4643 | -34.9831 | -27.358
17.4704 | -34.9836 | -27.3654
17.4765 | -34.9841 | -27.3728
17.4826 | -34.9846 | -27.3801
17.4887 | -34.9851 | -27.3874
17.4948 | -34.9856 | -27.3949
17.5009 | -34.9861 | -27.4023
17.507 | -34.9866 | -27.4098
17.5131 | -34.9871 | -27.4174
17.5192 | -34.9876 | -27.4246
17.5253 | -34.988 | -27.4319
17.5314 | -34.9885 | -27.4393
17.5375 | -34.989 | -27.4466
17.5436 | -34.9895 | -27.4542
17.5497 | -34.99 | -27.4616
17.5559 | -34.9905 | -27.4687
17.562 | -34.9911 | -27.4759
17.5681 | -34.9916 | -27.4833
17.5742 | -34.9921 | -27.4907
17.5803 | -34.9926 | -27.4982
17.5864 | -34.9931 | -27.5054
17.5925 | -34.9936 | -27.5129
17.5986 | -34.9941 | -27.5201
17.6047 | -34.9946 | -27.5276
17.6108 | -34.9951 | -27.5351
17.6169 | -34.9956 | -27.5424
17.623 | -34.9961 | -27.5496
17.6291 | -34.9966 | -27.5568
17.6352 | -34.9971 | -27.5642
17.6413 | -34.9976 | -27.5716
17.6474 | -34.9981 | -27.5791
17.6535 | -34.9986 | -27.5866
17.6596 | -34.9991 | -27.5936
17.6657 | -34.9996 | -27.6011
17.6718 | -35.0001 | -27.6084
17.6779 | -35.0006 | -27.6156
17.684 | -35.0011 | -27.6231
17.6901 | -35.0016 | -27.6306
17.6963 | -35.0021 | -27.6378
17.7024 | -35.0026 | -27.6451
17.7085 | -35.0031 | -27.6523
17.7146 | -35.0036 | -27.6596
17.7207 | -35.0041 | -27.6669
17.7268 | -35.0047 | -27.6741
17.7329 | -35.0051 | -27.6815
17.739 | -35.0056 | -27.6889
17.7451 | -35.0061 | -27.6962
17.7512 | -35.0066 | -27.7035
17.7573 | -35.0071 | -27.7109
17.7634 | -35.0076 | -27.7179
17.7695 | -35.0081 | -27.7252
17.7756 | -35.0086 | -27.7326
17.7817 | -35.0091 | -27.74
17.7878 | -35.0096 | -27.7475
17.7939 | -35.0101 | -27.7546
17.8 | -35.0105 | -27.7616
17.8061 | -35.0111 | -27.769
17.8122 | -35.0116 | -27.7764
17.8183 | -35.0121 | -27.7836
17.8244 | -35.0126 | -27.791
17.8305 | -35.013 | -27.7984
17.8366 | -35.0135 | -27.8054
17.8428 | -35.014 | -27.8127
17.8489 | -35.0145 | -27.82
17.855 | -35.015 | -27.8271
17.8611 | -35.0155 | -27.8343
17.8672 | -35.016 | -27.8413
17.8733 | -35.0165 | -27.8486
17.8794 | -35.017 | -27.856
17.8855 | -35.0175 | -27.8633
17.8916 | -35.0179 | -27.8706
17.8977 | -35.0184 | -27.8779
17.9038 | -35.0189 | -27.8849
17.9099 | -35.0194 | -27.8922
17.916 | -35.0199 | -27.8995
17.9221 | -35.0204 | -27.9069
17.9282 | -35.0209 | -27.914
17.9343 | -35.0214 | -27.9213
17.9404 | -35.0219 | -27.9283
17.9465 | -35.0224 | -27.9354
17.9526 | -35.0229 | -27.9429
17.9587 | -35.0234 | -27.9501
17.9648 | -35.0239 | -27.9574
17.9709 | -35.0244 | -27.9646
17.977 | -35.0249 | -27.9717
17.9832 | -35.0254 | -27.979
17.9893 | -35.0259 | -27.9864
17.9954 | -35.0263 | -27.9937
18.0015 | -35.0268 | -28.0011
18.0076 | -35.0273 | -28.0081
18.0137 | -35.0278 | -28.0155
18.0198 | -35.0283 | -28.0226
18.0259 | -35.0287 | -28.03
18.032 | -35.0292 | -28.0374
18.0381 | -35.0297 | -28.0447
18.0442 | -35.0302 | -28.0519
18.0503 | -35.0307 | -28.0592
18.0564 | -35.0312 | -28.0664
18.0625 | -35.0317 | -28.0738
18.0686 | -35.0321 | -28.081
18.0747 | -35.0326 | -28.0882
18.0808 | -35.0331 | -28.0955
18.0869 | -35.0336 | -28.1028
18.093 | -35.0341 | -28.1101
18.0991 | -35.0346 | -28.1175
18.1052 | -35.0351 | -28.1248
18.1113 | -35.0356 | -28.132
18.1174 | -35.0361 | -28.139
18.1236 | -35.0366 | -28.1462
18.1297 | -35.037 | -28.1534
18.1358 | -35.0375 | -28.1607
18.1419 | -35.038 | -28.1679
18.148 | -35.0385 | -28.1752
18.1541 | -35.039 | -28.1821
18.1602 | -35.0395 | -28.1894
18.1663 | -35.04 | -28.1965
18.1724 | -35.0405 | -28.2037
18.1785 | -35.0409 | -28.2109
18.1846 | -35.0414 | -28.2179
18.1907 | -35.0419 | -28.2251
18.1968 | -35.0423 | -28.2325
18.2029 | -35.0428 | -28.2397
18.209 | -35.0433 | -28.2469
18.2151 | -35.0438 | -28.2543
18.2212 | -35.0442 | -28.2615
18.2273 | -35.0447 | -28.2687
18.2334 | -35.0452 | -28.2758
18.2395 | -35.0457 | -28.2829
18.2456 | -35.0462 | -28.2903
18.2517 | -35.0467 | -28.2976
18.2578 | -35.0472 | -28.3048
18.2639 | -35.0476 | -28.3121
18.2701 | -35.0481 | -28.3193
18.2762 | -35.0486 | -28.3264
18.2823 | -35.0491 | -28.3335
18.2884 | -35.0496 | -28.3405
18.2945 | -35.0501 | -28.3475
18.3006 | -35.0506 | -28.3548
18.3067 | -35.0511 | -28.362
18.3128 | -35.0516 | -28.3693
18.3189 | -35.052 | -28.3765
18.325 | -35.0525 | -28.3834
18.3311 | -35.053 | -28.3906
18.3372 | -35.0535 | -28.3978
18.3433 | -35.054 | -28.4051
18.3494 | -35.0544 | -28.4123
18.3555 | -35.0549 | -28.4194
18.3616 | -35.0554 | -28.4264
18.3677 | -35.0559 | -28.4335
18.3738 | -35.0564 | -28.4407
18.3799 | -35.0569 | -28.4479
18.386 | -35.0573 | -28.4551
18.3921 | -35.0578 | -28.4624
18.3982 | -35.0583 | -28.4694
18.4043 | -35.0587 | -28.4766
18.4105 | -35.0592 | -28.4838
18.4166 | -35.0597 | -28.4911
18.4227 | -35.0602 | -28.4983
18.4288 | -35.0606 | -28.5053
18.4349 | -35.0611 | -28.5123
18.441 | -35.0616 | -28.5194
18.4471 | -35.062 | -28.5265
18.4532 | -35.0625 | -28.5336
18.4593 | -35.063 | -28.5408
18.4654 | -35.0635 | -28.5477
18.4715 | -35.0639 | -28.5549
18.4776 | -35.0644 | -28.5622
18.4837 | -35.0649 | -28.5694
18.4898 | -35.0653 | -28.5767
18.4959 | -35.0658 | -28.5839
18.502 | -35.0662 | -28.5909
18.5081 | -35.0667 | -28.5981
18.5142 | -35.0672 | -28.6052
18.5203 | -35.0676 | -28.6124
18.5264 | -35.0681 | -28.6194
18.5325 | -35.0686 | -28.6264
18.5386 | -35.0691 | -28.6335
18.5447 | -35.0695 | -28.6406
18.5508 | -35.07 | -28.6477
18.557 | -35.0705 | -28.6549
18.5631 | -35.0709 | -28.6622
18.5692 | -35.0714 | -28.6693
18.5753 | -35.0718 | -28.6762
18.5814 | -35.0723 | -28.6832
18.5875 | -35.0728 | -28.6903
18.5936 | -35.0733 | -28.6974
18.5997 | -35.0738 | -28.7047
18.6058 | -35.0742 | -28.7118
18.6119 | -35.0747 | -28.7189
18.618 | -35.0752 | -28.7263
18.6241 | -35.0756 | -28.7334
18.6302 | -35.0761 | -28.7407
18.6363 | -35.0766 | -28.7479
18.6424 | -35.077 | -28.755
18.6485 | -35.0775 | -28.7621
18.6546 | -35.078 | -28.7693
18.6607 | -35.0784 | -28.7765
18.6668 | -35.0789 | -28.7836
18.6729 | -35.0794 | -28.7909
18.679 | -35.0798 | -28.798
18.6851 | -35.0803 | -28.805
18.6912 | -35.0807 | -28.8121
18.6974 | -35.0812 | -28.8189
18.7035 | -35.0817 | -28.8259
18.7096 | -35.0822 | -28.8332
18.7157 | -35.0826 | -28.8403
18.7218 | -35.0831 | -28.8474
18.7279 | -35.0836 | -28.8546
18.734 | -35.084 | -28.8619
18.7401 | -35.0845 | -28.869
18.7462 | -35.085 | -28.8758
18.7523 | -35.0854 | -28.883
18.7584 | -35.0859 | -28.8901
18.7645 | -35.0864 | -28.8973
18.7706 | -35.0869 | -28.9043
18.7767 | -35.0873 | -28.9113
18.7828 | -35.0878 | -28.9181
18.7889 | -35.0883 | -28.9254
18.795 | -35.0887 | -28.9326
18.8011 | -35.0892 | -28.9398
18.8072 | -35.0897 | -28.947
18.8133 | -35.0901 | -28.9542
18.8194 | -35.0906 | -28.961
18.8255 | -35.0911 | -28.968
18.8316 | -35.0915 | -28.9752
18.8377 | -35.092 | -28.982
18.8439 | -35.0925 | -28.989
18.85 | -35.0929 | -28.9962
18.8561 | -35.0934 | -29.0033
18.8622 | -35.0938 | -29.0105
18.8683 | -35.0943 | -29.0174
18.8744 | -35.0948 | -29.0245
18.8805 | -35.0952 | -29.0318
18.8866 | -35.0957 | -29.0391
18.8927 | -35.0962 | -29.0459
18.8988 | -35.0966 | -29.0529
18.9049 | -35.0971 | -29.06
18.911 | -35.0976 | -29.0674
18.9171 | -35.0981 | -29.0744
18.9232 | -35.0986 | -29.0813
18.9293 | -35.099 | -29.0885
18.9354 | -35.0995 | -29.0957
18.9415 | -35.1 | -29.1028
18.9476 | -35.1004 | -29.1098
18.9537 | -35.1009 | -29.1169
18.9598 | -35.1014 | -29.1239
18.9659 | -35.1019 | -29.1309
18.972 | -35.1023 | -29.1379
18.9781 | -35.1027 | -29.145
18.9843 | -35.1032 | -29.1521
18.9904 | -35.1037 | -29.1591
18.9965 | -35.1041 | -29.1659
19.0026 | -35.1046 | -29.1732
19.0087 | -35.1051 | -29.1805
19.0148 | -35.1056 | -29.1875
19.0209 | -35.106 | -29.1946
19.027 | -35.1065 | -29.2018
19.0331 | -35.1069 | -29.2088
19.0392 | -35.1074 | -29.2159
19.0453 | -35.1079 | -29.2229
19.0514 | -35.1083 | -29.2299
19.0575 | -35.1088 | -29.2371
19.0636 | -35.1093 | -29.2441
19.0697 | -35.1097 | -29.251
19.0758 | -35.1102 | -29.258
19.0819 | -35.1107 | -29.265
19.088 | -35.1111 | -29.2722
19.0941 | -35.1116 | -29.2793
19.1002 | -35.112 | -29.2865
19.1063 | -35.1125 | -29.2936
19.1124 | -35.1129 | -29.3007
19.1185 | -35.1134 | -29.3075
19.1246 | -35.1139 | -29.3145
19.1308 | -35.1143 | -29.3216
19.1369 | -35.1148 | -29.3286
19.143 | -35.1153 | -29.3357
19.1491 | -35.1157 | -29.3426
19.1552 | -35.1162 | -29.3498
19.1613 | -35.1167 | -29.3569
19.1674 | -35.1171 | -29.3639
19.1735 | -35.1176 | -29.371
19.1796 | -35.118 | -29.3778
19.1857 | -35.1185 | -29.3851
19.1918 | -35.119 | -29.3922
19.1979 | -35.1195 | -29.3995
19.204 | -35.1199 | -29.4066
19.2101 | -35.1204 | -29.4137
19.2162 | -35.1208 | -29.4208
19.2223 | -35.1213 | -29.4278
19.2284 | -35.1217 | -29.4349
19.2345 | -35.1222 | -29.4423
19.2406 | -35.1226 | -29.4491
19.2467 | -35.1231 | -29.4564
19.2528 | -35.1235 | -29.4635
19.2589 | -35.1239 | -29.4703
19.265 | -35.1244 | -29.4772
19.2712 | -35.1249 | -29.4843
19.2773 | -35.1253 | -29.4912
19.2834 | -35.1257 | -29.4981
19.2895 | -35.1262 | -29.5051
19.2956 | -35.1266 | -29.5122
19.3017 | -35.1271 | -29.5192
19.3078 | -35.1275 | -29.5265
19.3139 | -35.128 | -29.5334
19.32 | -35.1284 | -29.5404
19.3261 | -35.1289 | -29.5475
19.3322 | -35.1293 | -29.5544
19.3383 | -35.1298 | -29.5616
19.3444 | -35.1302 | -29.5683
19.3505 | -35.1306 | -29.5753
19.3566 | -35.131 | -29.5824
19.3627 | -35.1315 | -29.5894
19.3688 | -35.1319 | -29.5962
19.3749 | -35.1324 | -29.6035
19.381 | -35.1328 | -29.6103
19.3871 | -35.1333 | -29.6173
19.3932 | -35.1337 | -29.6244
19.3993 | -35.1342 | -29.6314
19.4054 | -35.1347 | -29.6384
19.4115 | -35.1351 | -29.6456
19.4177 | -35.1356 | -29.6524
19.4238 | -35.136 | -29.6593
19.4299 | -35.1365 | -29.6661
19.436 | -35.1369 | -29.6732
19.4421 | -35.1374 | -29.6801
19.4482 | -35.1378 | -29.6871
19.4543 | -35.1383 | -29.6941
19.4604 | -35.1388 | -29.7011
19.4665 | -35.1392 | -29.708
19.4726 | -35.1396 | -29.7148
19.4787 | -35.1401 | -29.7218
19.4848 | -35.1405 | -29.7286
19.4909 | -35.141 | -29.7355
19.497 | -35.1414 | -29.7427
19.5031 | -35.1419 | -29.7498
19.5092 | -35.1423 | -29.7569
19.5153 | -35.1428 | -29.7639
19.5214 | -35.1432 | -29.7711
19.5275 | -35.1436 | -29.778
19.5336 | -35.1441 | -29.7849
19.5397 | -35.1446 | -29.7917
19.5458 | -35.145 | -29.7988
19.5519 | -35.1454 | -29.8057
19.5581 | -35.1458 | -29.8126
19.5642 | -35.1463 | -29.8197
19.5703 | -35.1467 | -29.8271
19.5764 | -35.1472 | -29.8343
19.5825 | -35.1476 | -29.8416
19.5886 | -35.1481 | -29.8486
19.5947 | -35.1485 | -29.8552
19.6008 | -35.1489 | -29.8622
19.6069 | -35.1494 | -29.8694
19.613 | -35.1498 | -29.8766
19.6191 | -35.1502 | -29.8835
19.6252 | -35.1507 | -29.8904
19.6313 | -35.1511 | -29.8976
19.6374 | -35.1515 | -29.9044
19.6435 | -35.152 | -29.9115
19.6496 | -35.1524 | -29.9185
19.6557 | -35.1528 | -29.9255
19.6618 | -35.1533 | -29.9325
19.6679 | -35.1538 | -29.9395
19.674 | -35.1542 | -29.9465
19.6801 | -35.1546 | -29.9535
19.6862 | -35.155 | -29.9604
19.6923 | -35.1555 | -29.9675
19.6984 | -35.1559 | -29.9742
19.7046 | -35.1564 | -29.9812
19.7107 | -35.1569 | -29.9881
19.7168 | -35.1573 | -29.9951
19.7229 | -35.1578 | -30.0021
19.729 | -35.1583 | -30.009
19.7351 | -35.1587 | -30.0157
19.7412 | -35.1592 | -30.0227
19.7473 | -35.1597 | -30.0295
19.7534 | -35.1601 | -30.0367
19.7595 | -35.1605 | -30.0439
19.7656 | -35.161 | -30.0509
19.7717 | -35.1614 | -30.0572
19.7778 | -35.1619 | -30.0643
19.7839 | -35.1623 | -30.0712
19.79 | -35.1628 | -30.0782
19.7961 | -35.1632 | -30.0853
19.8022 | -35.1636 | -30.0924
19.8083 | -35.1641 | -30.0994
19.8144 | -35.1646 | -30.1064
19.8205 | -35.165 | -30.1134
19.8266 | -35.1654 | -30.1201
19.8327 | -35.1659 | -30.127
19.8388 | -35.1663 | -30.1338
19.845 | -35.1668 | -30.1409
19.8511 | -35.1672 | -30.148
19.8572 | -35.1677 | -30.1548
19.8633 | -35.1681 | -30.1618
19.8694 | -35.1685 | -30.1687
19.8755 | -35.169 | -30.1755
19.8816 | -35.1694 | -30.1826
19.8877 | -35.1698 | -30.1897
19.8938 | -35.1702 | -30.1968
19.8999 | -35.1707 | -30.2039
19.906 | -35.1711 | -30.211
19.9121 | -35.1716 | -30.2179
19.9182 | -35.172 | -30.225
19.9243 | -35.1724 | -30.2321
19.9304 | -35.1729 | -30.2388
19.9365 | -35.1733 | -30.2458
19.9426 | -35.1738 | -30.253
19.9487 | -35.1742 | -30.2598
19.9548 | -35.1746 | -30.2669
19.9609 | -35.175 | -30.2738
19.967 | -35.1755 | -30.2806
19.9731 | -35.1759 | -30.2874
19.9792 | -35.1764 | -30.2943
19.9853 | -35.1768 | -30.3009
19.9915 | -35.1773 | -30.3078
19.9976 | -35.1777 | -30.3146
20.0037 | -35.1782 | -30.3217
20.0098 | -35.1786 | -30.3288
20.0159 | -35.1791 | -30.3358
20.022 | -35.1796 | -30.3424
20.0281 | -35.18 | -30.3494
20.0342 | -35.1804 | -30.3564
20.0403 | -35.1809 | -30.3634
20.0464 | -35.1813 | -30.3705
20.0525 | -35.1817 | -30.3774
20.0586 | -35.1822 | -30.384
20.0647 | -35.1826 | -30.3909
20.0708 | -35.1831 | -30.3979
20.0769 | -35.1835 | -30.4049
20.083 | -35.184 | -30.4118
20.0891 | -35.1844 | -30.4189
20.0952 | -35.1848 | -30.4257
20.1013 | -35.1852 | -30.4327
20.1074 | -35.1857 | -30.4395
20.1135 | -35.1862 | -30.4465
20.1196 | -35.1866 | -30.453
20.1257 | -35.187 | -30.4599
20.1319 | -35.1874 | -30.467
20.138 | -35.1879 | -30.4738
20.1441 | -35.1883 | -30.4805
20.1502 | -35.1888 | -30.4875
20.1563 | -35.1892 | -30.4945
20.1624 | -35.1896 | -30.5016
20.1685 | -35.1901 | -30.5087
20.1746 | -35.1906 | -30.5155
20.1807 | -35.191 | -30.5224
20.1868 | -35.1915 | -30.529
20.1929 | -35.1919 | -30.5364
20.199 | -35.1924 | -30.5435
20.2051 | -35.1929 | -30.5503
20.2112 | -35.1934 | -30.5571
20.2173 | -35.1938 | -30.564
20.2234 | -35.1942 | -30.5709
20.2295 | -35.1947 | -30.5779
20.2356 | -35.1952 | -30.5849
20.2417 | -35.1956 | -30.5918
20.2478 | -35.196 | -30.5984
20.2539 | -35.1965 | -30.6055
20.26 | -35.1969 | -30.6124
20.2661 | -35.1974 | -30.6197
20.2723 | -35.1978 | -30.6266
20.2784 | -35.1983 | -30.6336
20.2845 | -35.1987 | -30.6403
20.2906 | -35.1991 | -30.6472
20.2967 | -35.1996 | -30.654
20.3028 | -35.2 | -30.6611
20.3089 | -35.2004 | -30.668
20.315 | -35.2009 | -30.6749
20.3211 | -35.2014 | -30.682
20.3272 | -35.2018 | -30.6887
20.3333 | -35.2023 | -30.6959
20.3394 | -35.2028 | -30.7026
20.3455 | -35.2032 | -30.7094
20.3516 | -35.2036 | -30.7161
20.3577 | -35.204 | -30.7231
20.3638 | -35.2044 | -30.7301
20.3699 | -35.2049 | -30.7364
20.376 | -35.2053 | -30.7437
20.3821 | -35.2058 | -30.7506
20.3882 | -35.2062 | -30.7576
20.3943 | -35.2067 | -30.765
20.4004 | -35.207 | -30.7714
20.4065 | -35.2075 | -30.7784
20.4126 | -35.2079 | -30.7855
20.4188 | -35.2083 | -30.7925
20.4249 | -35.2087 | -30.7993
20.431 | -35.2091 | -30.8062
20.4371 | -35.2096 | -30.8129
20.4432 | -35.21 | -30.8197
20.4493 | -35.2104 | -30.8267
20.4554 | -35.2108 | -30.8335
20.4615 | -35.2113 | -30.8404
20.4676 | -35.2117 | -30.8475
20.4737 | -35.2122 | -30.8543
20.4798 | -35.2126 | -30.8613
20.4859 | -35.213 | -30.8683
20.492 | -35.2135 | -30.8749
20.4981 | -35.2139 | -30.8819
20.5042 | -35.2144 | -30.8888
20.5103 | -35.2148 | -30.8958
20.5164 | -35.2153 | -30.9025
20.5225 | -35.2157 | -30.9094
20.5286 | -35.2161 | -30.9165
20.5347 | -35.2165 | -30.9233
20.5408 | -35.217 | -30.9299
20.5469 | -35.2174 | -30.9368
20.553 | -35.2178 | -30.9438
20.5592 | -35.2182 | -30.9508
20.5653 | -35.2187 | -30.9578
20.5714 | -35.2191 | -30.9648
20.5775 | -35.2195 | -30.9715
20.5836 | -35.22 | -30.978
20.5897 | -35.2204 | -30.985
20.5958 | -35.2208 | -30.9919
20.6019 | -35.2213 | -30.999
20.608 | -35.2217 | -31.0057
20.6141 | -35.2223 | -31.0128
20.6202 | -35.2228 | -31.0197
20.6263 | -35.2233 | -31.0263
20.6324 | -35.2237 | -31.0332
20.6385 | -35.2241 | -31.04
20.6446 | -35.2245 | -31.0466
20.6507 | -35.225 | -31.0537
20.6568 | -35.2255 | -31.0604
20.6629 | -35.2259 | -31.0671
20.669 | -35.2263 | -31.0739
20.6751 | -35.2267 | -31.0805
20.6812 | -35.2271 | -31.0873
20.6873 | -35.2275 | -31.094
20.6934 | -35.2279 | -31.1007
20.6995 | -35.2284 | -31.1077
20.7057 | -35.2289 | -31.1147
20.7118 | -35.2293 | -31.1215
20.7179 | -35.2298 | -31.1279
20.724 | -35.2302 | -31.1346
20.7301 | -35.2306 | -31.1413
20.7362 | -35.2311 | -31.1485
20.7423 | -35.2315 | -31.1551
20.7484 | -35.2319 | -31.1616
20.7545 | -35.2325 | -31.1681
20.7606 | -35.2329 | -31.175
20.7667 | -35.2333 | -31.1817
20.7728 | -35.2338 | -31.1885
20.7789 | -35.2343 | -31.1951
20.785 | -35.2348 | -31.2022
20.7911 | -35.2351 | -31.2092
20.7972 | -35.2355 | -31.2165
20.8033 | -35.236 | -31.2233
20.8094 | -35.2365 | -31.2297
20.8155 | -35.2369 | -31.2368
20.8216 | -35.2374 | -31.2436
20.8277 | -35.2379 | -31.2498
20.8338 | -35.2383 | -31.2566
20.8399 | -35.2388 | -31.2631
20.8461 | -35.2391 | -31.2704
20.8522 | -35.2396 | -31.2776
20.8583 | -35.2401 | -31.2836
20.8644 | -35.2405 | -31.291
20.8705 | -35.2409 | -31.2979
20.8766 | -35.2413 | -31.305
20.8827 | -35.2417 | -31.3118
20.8888 | -35.2421 | -31.3186
20.8949 | -35.2426 | -31.325
20.901 | -35.243 | -31.3316
20.9071 | -35.2435 | -31.3382
20.9132 | -35.2439 | -31.3447
20.9193 | -35.2444 | -31.3516
20.9254 | -35.2448 | -31.3581
20.9315 | -35.2452 | -31.3652
20.9376 | -35.2457 | -31.3716
20.9437 | -35.2461 | -31.3787
20.9498 | -35.2466 | -31.386
20.9559 | -35.2471 | -31.3929
20.962 | -35.2474 | -31.3995
20.9681 | -35.2478 | -31.4065
20.9742 | -35.2483 | -31.4132
20.9803 | -35.2488 | -31.4204
20.9864 | -35.2492 | -31.4276
20.9926 | -35.2496 | -31.4339
20.9987 | -35.2501 | -31.441
21.0048 | -35.2506 | -31.4483
21.0109 | -35.251 | -31.455
21.017 | -35.2515 | -31.4619
21.0231 | -35.2519 | -31.469
21.0292 | -35.2524 | -31.4753
21.0353 | -35.2528 | -31.4823
21.0414 | -35.2531 | -31.4892
21.0475 | -35.2535 | -31.4958
21.0536 | -35.254 | -31.5025
21.0597 | -35.2544 | -31.5095
21.0658 | -35.2548 | -31.5164
21.0719 | -35.2552 | -31.5235
21.078 | -35.2557 | -31.5299
21.0841 | -35.2562 | -31.5367
21.0902 | -35.2566 | -31.5432
21.0963 | -35.2569 | -31.5497
21.1024 | -35.2573 | -31.5561
21.1085 | -35.2577 | -31.5629
21.1146 | -35.2582 | -31.5693
21.1207 | -35.2587 | -31.576
21.1268 | -35.259 | -31.5829
21.133 | -35.2595 | -31.5895
21.1391 | -35.26 | -31.5955
21.1452 | -35.2603 | -31.602
21.1513 | -35.2607 | -31.6085
21.1574 | -35.2611 | -31.6147
21.1635 | -35.2616 | -31.6215
21.1696 | -35.262 | -31.6284
21.1757 | -35.2625 | -31.6349
21.1818 | -35.263 | -31.6415
21.1879 | -35.2635 | -31.6479
21.194 | -35.2639 | -31.6545
21.2001 | -35.2644 | -31.6614
21.2062 | -35.2648 | -31.6685
21.2123 | -35.2653 | -31.6752
21.2184 | -35.2657 | -31.6816
21.2245 | -35.2661 | -31.6885
21.2306 | -35.2665 | -31.6953
21.2367 | -35.2669 | -31.7023
21.2428 | -35.2673 | -31.7088
21.2489 | -35.2677 | -31.7153
21.255 | -35.2681 | -31.7222
21.2611 | -35.2686 | -31.7285
21.2672 | -35.269 | -31.735
21.2733 | -35.2694 | -31.742
21.2795 | -35.2698 | -31.7489
21.2856 | -35.2703 | -31.7558
21.2917 | -35.2707 | -31.7626
21.2978 | -35.2712 | -31.7693
21.3039 | -35.2716 | -31.776
21.31 | -35.2721 | -31.7828
21.3161 | -35.2725 | -31.7891
21.3222 | -35.2729 | -31.796
21.3283 | -35.2733 | -31.8027
21.3344 | -35.2738 | -31.8094
21.3405 | -35.2742 | -31.8167
21.3466 | -35.2746 | -31.8233
21.3527 | -35.2751 | -31.8302
21.3588 | -35.2756 | -31.8365
21.3649 | -35.276 | -31.8427
21.371 | -35.2764 | -31.8495
21.3771 | -35.2769 | -31.856
21.3832 | -35.2772 | -31.8628
21.3893 | -35.2777 | -31.8693
21.3954 | -35.2781 | -31.8764
21.4015 | -35.2785 | -31.8833
21.4076 | -35.2789 | -31.8899
21.4137 | -35.2794 | -31.8967
21.4199 | -35.2799 | -31.9032
21.426 | -35.2803 | -31.9099
21.4321 | -35.2807 | -31.9166
21.4382 | -35.2812 | -31.9232
21.4443 | -35.2817 | -31.9297
21.4504 | -35.2821 | -31.9366
21.4565 | -35.2825 | -31.9426
21.4626 | -35.2829 | -31.9495
21.4687 | -35.2834 | -31.9562
21.4748 | -35.2839 | -31.9627
21.4809 | -35.2843 | -31.9689
21.487 | -35.2847 | -31.9755
21.4931 | -35.2851 | -31.9823
21.4992 | -35.2855 | -31.9891
21.5053 | -35.286 | -31.9958
21.5114 | -35.2865 | -32.0021
21.5175 | -35.2869 | -32.0085
21.5236 | -35.2873 | -32.0151
21.5297 | -35.2877 | -32.0218
21.5358 | -35.2881 | -32.0281
21.5419 | -35.2885 | -32.035
21.548 | -35.2889 | -32.0414
21.5541 | -35.2893 | -32.0483
21.5602 | -35.2898 | -32.0545
21.5664 | -35.2902 | -32.0611
21.5725 | -35.2907 | -32.0672
21.5786 | -35.2912 | -32.0739
21.5847 | -35.2916 | -32.0803
21.5908 | -35.292 | -32.0869
21.5969 | -35.2924 | -32.0928
21.603 | -35.2928 | -32.0999
21.6091 | -35.2934 | -32.1071
21.6152 | -35.2938 | -32.1134
21.6213 | -35.2942 | -32.1201
21.6274 | -35.2946 | -32.1264
21.6335 | -35.2951 | -32.1331
21.6396 | -35.2955 | -32.1398
21.6457 | -35.296 | -32.1462
21.6518 | -35.2963 | -32.1525
21.6579 | -35.2968 | -32.159
21.664 | -35.2973 | -32.1658
21.6701 | -35.2977 | -32.1729
21.6762 | -35.2981 | -32.1795
21.6823 | -35.2985 | -32.1857
21.6884 | -35.299 | -32.1925
21.6945 | -35.2994 | -32.1988
21.7006 | -35.2998 | -32.2054
21.7068 | -35.3003 | -32.2121
21.7129 | -35.3007 | -32.2189
21.719 | -35.3012 | -32.2253
21.7251 | -35.3017 | -32.2316
21.7312 | -35.302 | -32.2383
21.7373 | -35.3024 | -32.2441
21.7434 | -35.3029 | -32.2505
21.7495 | -35.3034 | -32.2571
21.7556 | -35.3038 | -32.2634
21.7617 | -35.3043 | -32.2695
21.7678 | -35.3047 | -32.2762
21.7739 | -35.3051 | -32.2819
21.78 | -35.3057 | -32.2885
21.7861 | -35.3062 | -32.2954
21.7922 | -35.3067 | -32.3019
21.7983 | -35.3071 | -32.308
21.8044 | -35.3075 | -32.3145
21.8105 | -35.308 | -32.3208
21.8166 | -35.3084 | -32.3272
21.8227 | -35.3088 | -32.3335
21.8288 | -35.3093 | -32.3396
21.8349 | -35.3096 | -32.3467
21.841 | -35.3101 | -32.3532
21.8471 | -35.3105 | -32.3597
21.8533 | -35.3111 | -32.3659
21.8594 | -35.3115 | -32.3725
21.8655 | -35.312 | -32.3792
21.8716 | -35.3125 | -32.3855
21.8777 | -35.313 | -32.3922
21.8838 | -35.3134 | -32.3985
21.8899 | -35.3138 | -32.4053
21.896 | -35.3143 | -32.4121
21.9021 | -35.3146 | -32.4188
21.9082 | -35.3151 | -32.4256
21.9143 | -35.3156 | -32.4327
21.9204 | -35.3159 | -32.439
21.9265 | -35.3164 | -32.4453
21.9326 | -35.3168 | -32.4518
21.9387 | -35.3172 | -32.4582
21.9448 | -35.3177 | -32.4648
21.9509 | -35.3181 | -32.4714
21.957 | -35.3186 | -32.478
21.9631 | -35.319 | -32.4844
21.9692 | -35.3194 | -32.4908
21.9753 | -35.3198 | -32.4975
21.9814 | -35.3202 | -32.5038
21.9875 | -35.3206 | -32.5099
21.9937 | -35.3211 | -32.5167
21.9998 | -35.3215 | -32.5233
22.0059 | -35.3219 | -32.5297
22.012 | -35.3224 | -32.5363
22.0181 | -35.3229 | -32.5424
22.0242 | -35.3233 | -32.5483
22.0303 | -35.3237 | -32.5547
22.0364 | -35.3242 | -32.5613
22.0425 | -35.3247 | -32.568
22.0486 | -35.3252 | -32.5747
22.0547 | -35.3256 | -32.581
22.0608 | -35.3261 | -32.5876
22.0669 | -35.3265 | -32.5939
22.073 | -35.327 | -32.6
22.0791 | -35.3274 | -32.6062
22.0852 | -35.3278 | -32.6127
22.0913 | -35.3282 | -32.6187
22.0974 | -35.3286 | -32.6249
22.1035 | -35.3291 | -32.6318
22.1096 | -35.3295 | -32.6382
22.1157 | -35.3299 | -32.6453
22.1218 | -35.3304 | -32.6515
22.1279 | -35.3308 | -32.6576
22.134 | -35.3312 | -32.6642
22.1402 | -35.3317 | -32.6708
22.1463 | -35.3321 | -32.6773
22.1524 | -35.3325 | -32.6833
22.1585 | -35.3329 | -32.6894
22.1646 | -35.3335 | -32.6956
22.1707 | -35.3339 | -32.7018
22.1768 | -35.3344 | -32.7083
22.1829 | -35.3348 | -32.7149
22.189 | -35.3353 | -32.7216
22.1951 | -35.3357 | -32.7277
22.2012 | -35.3362 | -32.7344
22.2073 | -35.3366 | -32.741
22.2134 | -35.337 | -32.7475
22.2195 | -35.3374 | -32.7538
22.2256 | -35.3379 | -32.76
22.2317 | -35.3384 | -32.7656
22.2378 | -35.3388 | -32.7721
22.2439 | -35.3393 | -32.7786
22.25 | -35.3397 | -32.785
22.2561 | -35.3401 | -32.7918
22.2622 | -35.3406 | -32.798
22.2683 | -35.3411 | -32.8044
22.2744 | -35.3415 | -32.8112
22.2806 | -35.3419 | -32.8181
22.2867 | -35.3424 | -32.8243
22.2928 | -35.3428 | -32.831
22.2989 | -35.3433 | -32.8374
22.305 | -35.3437 | -32.8438
22.3111 | -35.3441 | -32.8504
22.3172 | -35.3446 | -32.8569
22.3233 | -35.3451 | -32.8635
22.3294 | -35.3455 | -32.87
22.3355 | -35.346 | -32.8761
22.3416 | -35.3464 | -32.8827
22.3477 | -35.3469 | -32.8893
22.3538 | -35.3473 | -32.8952
22.3599 | -35.3477 | -32.9016
22.366 | -35.3481 | -32.9079
22.3721 | -35.3484 | -32.9137
22.3782 | -35.3489 | -32.9198
22.3843 | -35.3494 | -32.9267
22.3904 | -35.3498 | -32.9328
22.3965 | -35.3503 | -32.9394
22.4026 | -35.3506 | -32.9455
22.4087 | -35.351 | -32.9519
22.4148 | -35.3515 | -32.9578
22.4209 | -35.352 | -32.9644
22.4271 | -35.3524 | -32.9707
22.4332 | -35.3529 | -32.9774
22.4393 | -35.3532 | -32.9839
22.4454 | -35.3536 | -32.9902
22.4515 | -35.3541 | -32.9965
22.4576 | -35.3545 | -33.0028
22.4637 | -35.3549 | -33.0093
22.4698 | -35.3553 | -33.0161
22.4759 | -35.3558 | -33.0225
22.482 | -35.3563 | -33.0291
22.4881 | -35.3568 | -33.0353
22.4942 | -35.3572 | -33.0417
22.5003 | -35.3576 | -33.0482
22.5064 | -35.3581 | -33.0545
22.5125 | -35.3584 | -33.0604
22.5186 | -35.3588 | -33.067
22.5247 | -35.3592 | -33.0727
22.5308 | -35.3597 | -33.0791
22.5369 | -35.3602 | -33.0856
22.543 | -35.3606 | -33.0919
22.5491 | -35.3611 | -33.098
22.5552 | -35.3615 | -33.1048
22.5613 | -35.362 | -33.1112
22.5675 | -35.3624 | -33.1171
22.5736 | -35.3629 | -33.1235
22.5797 | -35.3633 | -33.1303
22.5858 | -35.3638 | -33.1362
22.5919 | -35.3642 | -33.1427
22.598 | -35.3647 | -33.1491
22.6041 | -35.3651 | -33.1556
22.6102 | -35.3656 | -33.1619
22.6163 | -35.366 | -33.1679
22.6224 | -35.3664 | -33.174
22.6285 | -35.3668 | -33.1803
22.6346 | -35.3672 | -33.1868
22.6407 | -35.3676 | -33.1932
22.6468 | -35.3681 | -33.1998
22.6529 | -35.3684 | -33.2059
22.659 | -35.3688 | -33.2124
22.6651 | -35.3693 | -33.2192
22.6712 | -35.3697 | -33.2258
22.6773 | -35.3701 | -33.2317
22.6834 | -35.3705 | -33.2386
22.6895 | -35.3711 | -33.2451
22.6956 | -35.3715 | -33.252
22.7017 | -35.3719 | -33.2579
22.7079 | -35.3724 | -33.2642
22.714 | -35.3727 | -33.271
22.7201 | -35.3731 | -33.2771
22.7262 | -35.3737 | -33.2827
22.7323 | -35.3741 | -33.2892
22.7384 | -35.3746 | -33.2953
22.7445 | -35.3749 | -33.302
22.7506 | -35.3754 | -33.3084
22.7567 | -35.3757 | -33.3144
22.7628 | -35.3762 | -33.3204
22.7689 | -35.3766 | -33.3265
22.775 | -35.377 | -33.3327
22.7811 | -35.3775 | -33.3387
22.7872 | -35.3779 | -33.3449
22.7933 | -35.3783 | -33.3508
22.7994 | -35.3788 | -33.3575
22.8055 | -35.3792 | -33.3638
22.8116 | -35.3796 | -33.3703
22.8177 | -35.3801 | -33.3764
22.8238 | -35.3806 | -33.3826
22.8299 | -35.3812 | -33.3885
22.836 | -35.3817 | -33.3949
22.8421 | -35.3821 | -33.4016
22.8482 | -35.3825 | -33.408
22.8544 | -35.383 | -33.414
22.8605 | -35.3834 | -33.4201
22.8666 | -35.3838 | -33.4265
22.8727 | -35.3842 | -33.4327
22.8788 | -35.3847 | -33.4389
22.8849 | -35.3851 | -33.4452
22.891 | -35.3855 | -33.4517
22.8971 | -35.386 | -33.4579
22.9032 | -35.3864 | -33.464
22.9093 | -35.3869 | -33.4702
22.9154 | -35.3874 | -33.4765
22.9215 | -35.3878 | -33.4825
22.9276 | -35.3882 | -33.4888
22.9337 | -35.3885 | -33.4949
22.9398 | -35.3889 | -33.5009
22.9459 | -35.3894 | -33.5071
22.952 | -35.3898 | -33.513
22.9581 | -35.3902 | -33.5197
22.9642 | -35.3907 | -33.5259
22.9703 | -35.3911 | -33.5324
22.9764 | -35.3916 | -33.5388
22.9825 | -35.3919 | -33.5449
22.9886 | -35.3923 | -33.5512
22.9948 | -35.3927 | -33.5574
23.0009 | -35.3931 | -33.5636
23.007 | -35.3936 | -33.5699
23.0131 | -35.3941 | -33.5759
23.0192 | -35.3944 | -33.5819
23.0253 | -35.3949 | -33.5879
23.0314 | -35.3953 | -33.5945
23.0375 | -35.3957 | -33.6006
23.0436 | -35.3961 | -33.6067
23.0497 | -35.3966 | -33.6129
23.0558 | -35.397 | -33.6186
23.0619 | -35.3974 | -33.625
23.068 | -35.3978 | -33.6309
23.0741 | -35.3983 | -33.6368
23.0802 | -35.3986 | -33.6429
23.0863 | -35.3991 | -33.6492
23.0924 | -35.3995 | -33.6554
23.0985 | -35.4 | -33.6613
23.1046 | -35.4004 | -33.6678
23.1107 | -35.4008 | -33.6741
23.1168 | -35.4013 | -33.6801
23.1229 | -35.4017 | -33.6864
23.129 | -35.4022 | -33.6929
23.1351 | -35.4026 | -33.6987
23.1413 | -35.403 | -33.7051
23.1474 | -35.4034 | -33.7119
23.1535 | -35.4039 | -33.7183
23.1596 | -35.4043 | -33.7247
23.1657 | -35.4048 | -33.7308
23.1718 | -35.4052 | -33.7367
23.1779 | -35.4057 | -33.7427
23.184 | -35.4061 | -33.7489
23.1901 | -35.4066 | -33.7551
23.1962 | -35.407 | -33.7614
23.2023 | -35.4074 | -33.7674
23.2084 | -35.4079 | -33.7736
23.2145 | -35.4083 | -33.7793
23.2206 | -35.4088 | -33.7855
23.2267 | -35.4093 | -33.7913
23.2328 | -35.4097 | -33.7974
23.2389 | -35.4102 | -33.8035
23.245 | -35.4106 | -33.8098
23.2511 | -35.4111 | -33.8155
23.2572 | -35.4115 | -33.8217
23.2633 | -35.4119 | -33.8278
23.2694 | -35.4123 | -33.8337
23.2755 | -35.4126 | -33.84
23.2817 | -35.4131 | -33.8461
23.2878 | -35.4135 | -33.8523
23.2939 | -35.414 | -33.8581
23.3 | -35.4144 | -33.8645
23.3061 | -35.4148 | -33.8705
23.3122 | -35.4152 | -33.8769
23.3183 | -35.4157 | -33.8828
23.3244 | -35.4161 | -33.8893
23.3305 | -35.4165 | -33.8958
23.3366 | -35.417 | -33.9017
23.3427 | -35.4175 | -33.9077
23.3488 | -35.4179 | -33.914
23.3549 | -35.4183 | -33.9209
23.361 | -35.4186 | -33.9262
23.3671 | -35.4191 | -33.9328
23.3732 | -35.4196 | -33.939
23.3793 | -35.42 | -33.9448
23.3854 | -35.4205 | -33.9511
23.3915 | -35.421 | -33.9575
23.3976 | -35.4215 | -33.9638
23.4037 | -35.4219 | -33.9697
23.4098 | -35.4224 | -33.9755
23.4159 | -35.4228 | -33.9818
23.422 | -35.4232 | -33.988
23.4282 | -35.4237 | -33.9941
23.4343 | -35.4242 | -34.0005
23.4404 | -35.4246 | -34.007
23.4465 | -35.4251 | -34.0133
23.4526 | -35.4255 | -34.0193
23.4587 | -35.4259 | -34.0251
23.4648 | -35.4263 | -34.0313
23.4709 | -35.4267 | -34.0373
23.477 | -35.4272 | -34.0435
23.4831 | -35.4276 | -34.0496
23.4892 | -35.428 | -34.0555
23.4953 | -35.4285 | -34.0619
23.5014 | -35.429 | -34.0681
23.5075 | -35.4294 | -34.0743
23.5136 | -35.4299 | -34.0805
23.5197 | -35.4303 | -34.0865
23.5258 | -35.4307 | -34.0924
23.5319 | -35.4312 | -34.0981
23.538 | -35.4317 | -34.1045
23.5441 | -35.4321 | -34.1106
23.5502 | -35.4325 | -34.1166
23.5563 | -35.4329 | -34.1226
23.5624 | -35.4334 | -34.1289
23.5686 | -35.4337 | -34.1351
23.5747 | -35.4342 | -34.1407
23.5808 | -35.4347 | -34.1464
23.5869 | -35.4351 | -34.1524
23.593 | -35.4355 | -34.1588
23.5991 | -35.4359 | -34.1648
23.6052 | -35.4363 | -34.1709
23.6113 | -35.4367 | -34.1772
23.6174 | -35.4372 | -34.1836
23.6235 | -35.4376 | -34.1899
23.6296 | -35.4381 | -34.1965
23.6357 | -35.4385 | -34.2028
23.6418 | -35.4388 | -34.209
23.6479 | -35.4392 | -34.2145
23.654 | -35.4397 | -34.2207
23.6601 | -35.44 | -34.2269
23.6662 | -35.4405 | -34.2327
23.6723 | -35.4409 | -34.2386
23.6784 | -35.4414 | -34.2445
23.6845 | -35.4418 | -34.2506
23.6906 | -35.4422 | -34.2566
23.6967 | -35.4426 | -34.2628
23.7028 | -35.443 | -34.2689
23.7089 | -35.4434 | -34.2746
23.7151 | -35.4438 | -34.2813
23.7212 | -35.4443 | -34.2874
23.7273 | -35.4447 | -34.2934
23.7334 | -35.4451 | -34.2995
23.7395 | -35.4455 | -34.3056
23.7456 | -35.4459 | -34.3111
23.7517 | -35.4463 | -34.3174
23.7578 | -35.4467 | -34.3234
23.7639 | -35.4471 | -34.3294
23.77 | -35.4475 | -34.3358
23.7761 | -35.4479 | -34.3421
23.7822 | -35.4484 | -34.3481
23.7883 | -35.4488 | -34.3541
23.7944 | -35.4493 | -34.3601
23.8005 | -35.4497 | -34.3665
23.8066 | -35.4502 | -34.3728
23.8127 | -35.4506 | -34.3782
23.8188 | -35.451 | -34.3844
23.8249 | -35.4515 | -34.3907
23.831 | -35.4519 | -34.3971
23.8371 | -35.4523 | -34.4031
23.8432 | -35.4528 | -34.4092
23.8493 | -35.4531 | -34.4153
23.8555 | -35.4535 | -34.4211
23.8616 | -35.4539 | -34.427
23.8677 | -35.4543 | -34.433
23.8738 | -35.4548 | -34.4396
23.8799 | -35.4551 | -34.4455
23.886 | -35.4556 | -34.4506
23.8921 | -35.456 | -34.4571
23.8982 | -35.4564 | -34.463
23.9043 | -35.4568 | -34.4691
23.9104 | -35.4572 | -34.4753
23.9165 | -35.4577 | -34.4811
23.9226 | -35.458 | -34.4871
23.9287 | -35.4585 | -34.4932
23.9348 | -35.4589 | -34.4991
23.9409 | -35.4593 | -34.5049
23.947 | -35.4597 | -34.5109
23.9531 | -35.4602 | -34.5163
23.9592 | -35.4607 | -34.5221
23.9653 | -35.4611 | -34.5284
23.9714 | -35.4615 | -34.5343
23.9775 | -35.462 | -34.5408
23.9836 | -35.4624 | -34.5471
23.9897 | -35.4628 | -34.5524
23.9958 | -35.4633 | -34.5583
24.002 | -35.4636 | -34.5641
24.0081 | -35.4641 | -34.5705
24.0142 | -35.4645 | -34.5764
24.0203 | -35.4649 | -34.5825
24.0264 | -35.4653 | -34.5883
24.0325 | -35.4657 | -34.5942
24.0386 | -35.4661 | -34.6007
24.0447 | -35.4664 | -34.6068
24.0508 | -35.4667 | -34.6132
24.0569 | -35.4672 | -34.6191
24.063 | -35.4677 | -34.6254
24.0691 | -35.4681 | -34.6316
24.0752 | -35.4685 | -34.6377
24.0813 | -35.4689 | -34.644
24.0874 | -35.4694 | -34.6501
24.0935 | -35.4698 | -34.6558
24.0996 | -35.4702 | -34.6612
24.1057 | -35.4705 | -34.6674
24.1118 | -35.471 | -34.6736
24.1179 | -35.4713 | -34.6799
24.124 | -35.4717 | -34.6853
24.1301 | -35.4721 | -34.6909
24.1362 | -35.4725 | -34.6968
24.1424 | -35.4729 | -34.7029
24.1485 | -35.4733 | -34.7088
24.1546 | -35.4737 | -34.7148
24.1607 | -35.4741 | -34.7209
24.1668 | -35.4745 | -34.7268
24.1729 | -35.4749 | -34.7328
24.179 | -35.4754 | -34.7388
24.1851 | -35.4758 | -34.7451
24.1912 | -35.4761 | -34.7513
24.1973 | -35.4766 | -34.7574
24.2034 | -35.477 | -34.7632
24.2095 | -35.4774 | -34.7695
24.2156 | -35.4777 | -34.7755
24.2217 | -35.4782 | -34.7816
24.2278 | -35.4786 | -34.7877
24.2339 | -35.479 | -34.7938
24.24 | -35.4793 | -34.7992
24.2461 | -35.4798 | -34.8051
24.2522 | -35.4801 | -34.8112
24.2583 | -35.4805 | -34.8172
24.2644 | -35.481 | -34.8233
24.2705 | -35.4814 | -34.8291
24.2766 | -35.4819 | -34.8347
24.2827 | -35.4823 | -34.8406
24.2889 | -35.4827 | -34.8466
24.295 | -35.4831 | -34.8524
24.3011 | -35.4835 | -34.8582
24.3072 | -35.484 | -34.8638
24.3133 | -35.4843 | -34.8698
24.3194 | -35.4847 | -34.8755
24.3255 | -35.4852 | -34.8817
24.3316 | -35.4856 | -34.888
24.3377 | -35.4861 | -34.894
24.3438 | -35.4864 | -34.8994
24.3499 | -35.4868 | -34.905
24.356 | -35.4872 | -34.9109
24.3621 | -35.4876 | -34.9164
24.3682 | -35.488 | -34.9224
24.3743 | -35.4884 | -34.9283
24.3804 | -35.4889 | -34.934
24.3865 | -35.4893 | -34.9397
24.3926 | -35.4897 | -34.945
24.3987 | -35.49 | -34.9511
24.4048 | -35.4905 | -34.9567
24.4109 | -35.491 | -34.9625
24.417 | -35.4915 | -34.9687
24.4231 | -35.4918 | -34.9745
24.4293 | -35.4923 | -34.9802
24.4354 | -35.4927 | -34.9863
24.4415 | -35.4931 | -34.9923
24.4476 | -35.4935 | -34.9975
24.4537 | -35.4939 | -35.0036
24.4598 | -35.4943 | -35.0095
24.4659 | -35.4947 | -35.0156
24.472 | -35.4951 | -35.0211
24.4781 | -35.4955 | -35.0273
24.4842 | -35.4959 | -35.033
24.4903 | -35.4962 | -35.0393
24.4964 | -35.4967 | -35.0454
24.5025 | -35.497 | -35.0515
24.5086 | -35.4974 | -35.0573
24.5147 | -35.4977 | -35.0632
24.5208 | -35.4981 | -35.0689
24.5269 | -35.4985 | -35.0748
24.533 | -35.4989 | -35.0809
24.5391 | -35.4993 | -35.0865
24.5452 | -35.4998 | -35.0921
24.5513 | -35.5002 | -35.0975
24.5574 | -35.5006 | -35.1031
24.5635 | -35.501 | -35.1094
24.5696 | -35.5014 | -35.1153
24.5758 | -35.5018 | -35.1217
24.5819 | -35.5022 | -35.1272
24.588 | -35.5026 | -35.1332
24.5941 | -35.503 | -35.1396
24.6002 | -35.5034 | -35.1454
24.6063 | -35.5038 | -35.1515
24.6124 | -35.5042 | -35.1565
24.6185 | -35.5046 | -35.1625
24.6246 | -35.505 | -35.1683
24.6307 | -35.5053 | -35.1743
24.6368 | -35.5057 | -35.1806
24.6429 | -35.5061 | -35.1861
24.649 | -35.5065 | -35.1919
24.6551 | -35.5069 | -35.1982
24.6612 | -35.5073 | -35.2036
24.6673 | -35.5078 | -35.2095
24.6734 | -35.5082 | -35.2155
24.6795 | -35.5085 | -35.2215
24.6856 | -35.5089 | -35.2275
24.6917 | -35.5094 | -35.2333
24.6978 | -35.5099 | -35.2387
24.7039 | -35.5103 | -35.2444
24.71 | -35.5107 | -35.2504
24.7162 | -35.5111 | -35.2563
24.7223 | -35.5115 | -35.262
24.7284 | -35.5119 | -35.2668
24.7345 | -35.5122 | -35.2724
24.7406 | -35.5127 | -35.2779
24.7467 | -35.5131 | -35.2841
24.7528 | -35.5135 | -35.2901
24.7589 | -35.514 | -35.2958
24.765 | -35.5143 | -35.3015
24.7711 | -35.5147 | -35.3075
24.7772 | -35.515 | -35.3139
24.7833 | -35.5154 | -35.3197
24.7894 | -35.5158 | -35.3255
24.7955 | -35.5163 | -35.3313
24.8016 | -35.5167 | -35.3375
24.8077 | -35.5171 | -35.3435
24.8138 | -35.5175 | -35.3499
24.8199 | -35.5179 | -35.3555
24.826 | -35.5183 | -35.3612
24.8321 | -35.5187 | -35.3669
24.8382 | -35.5193 | -35.3722
24.8443 | -35.5198 | -35.3778
24.8504 | -35.5202 | -35.3841
24.8565 | -35.5205 | -35.3898
24.8627 | -35.5209 | -35.3957
24.8688 | -35.5213 | -35.4012
24.8749 | -35.5217 | -35.4068
24.881 | -35.522 | -35.4125
24.8871 | -35.5224 | -35.4185
24.8932 | -35.5228 | -35.4243
24.8993 | -35.5232 | -35.4302
24.9054 | -35.5236 | -35.4361
24.9115 | -35.524 | -35.4419
24.9176 | -35.5245 | -35.4475
24.9237 | -35.5249 | -35.4534
24.9298 | -35.5252 | -35.4591
24.9359 | -35.5256 | -35.4647
24.942 | -35.5261 | -35.4704
24.9481 | -35.5265 | -35.4766
24.9542 | -35.5268 | -35.4825
24.9603 | -35.5271 | -35.4879
24.9664 | -35.5276 | -35.4938
24.9725 | -35.5279 | -35.4995
24.9786 | -35.5282 | -35.5043
24.9847 | -35.5287 | -35.5102
24.9908 | -35.5291 | -35.516
24.9969 | -35.5294 | -35.522
25.0031 | -35.5298 | -35.5272
25.0092 | -35.5303 | -35.5333
25.0153 | -35.5307 | -35.5392
25.0214 | -35.5312 | -35.5451
25.0275 | -35.5316 | -35.5502
25.0336 | -35.532 | -35.5559
25.0397 | -35.5324 | -35.5615
25.0458 | -35.5329 | -35.5671
25.0519 | -35.5332 | -35.5732
25.058 | -35.5336 | -35.5792
25.0641 | -35.534 | -35.5842
25.0702 | -35.5343 | -35.5899
25.0763 | -35.5347 | -35.5955
25.0824 | -35.5351 | -35.6011
25.0885 | -35.5355 | -35.6068
25.0946 | -35.5358 | -35.6126
25.1007 | -35.5362 | -35.6184
25.1068 | -35.5366 | -35.624
25.1129 | -35.5371 | -35.6296
25.119 | -35.5375 | -35.635
25.1251 | -35.5379 | -35.6406
25.1312 | -35.5384 | -35.6462
25.1373 | -35.5388 | -35.6521
25.1435 | -35.5392 | -35.6578
25.1496 | -35.5396 | -35.6639
25.1557 | -35.54 | -35.6697
25.1618 | -35.5403 | -35.6755
25.1679 | -35.5407 | -35.6814
25.174 | -35.5411 | -35.687
25.1801 | -35.5414 | -35.6929
25.1862 | -35.5417 | -35.6984
25.1923 | -35.5422 | -35.7045
25.1984 | -35.5426 | -35.71
25.2045 | -35.543 | -35.7156
25.2106 | -35.5434 | -35.7209
25.2167 | -35.5437 | -35.7266
25.2228 | -35.5442 | -35.7327
25.2289 | -35.5445 | -35.7384
25.235 | -35.545 | -35.7443
25.2411 | -35.5453 | -35.7499
25.2472 | -35.5457 | -35.7556
25.2533 | -35.5461 | -35.7614
25.2594 | -35.5464 | -35.7671
25.2655 | -35.5468 | -35.7728
25.2716 | -35.5471 | -35.7784
25.2777 | -35.5475 | -35.7839
25.2838 | -35.5479 | -35.7895
25.29 | -35.5483 | -35.7956
25.2961 | -35.5486 | -35.8012
25.3022 | -35.549 | -35.8066
25.3083 | -35.5495 | -35.812
25.3144 | -35.5499 | -35.8177
25.3205 | -35.5503 | -35.8236
25.3266 | -35.5507 | -35.8286
25.3327 | -35.5511 | -35.8345
25.3388 | -35.5515 | -35.8405
25.3449 | -35.5519 | -35.8457
25.351 | -35.5522 | -35.8516
25.3571 | -35.5526 | -35.8571
25.3632 | -35.5531 | -35.8628
25.3693 | -35.5535 | -35.8686
25.3754 | -35.5539 | -35.8744
25.3815 | -35.5543 | -35.8798
25.3876 | -35.5547 | -35.8854
25.3937 | -35.5551 | -35.8909
25.3998 | -35.5554 | -35.8966
25.4059 | -35.5558 | -35.9026
25.412 | -35.5562 | -35.9082
25.4181 | -35.5566 | -35.9141
25.4242 | -35.557 | -35.9199
25.4304 | -35.5573 | -35.9254
25.4365 | -35.5578 | -35.9308
25.4426 | -35.5581 | -35.9363
25.4487 | -35.5584 | -35.9422
25.4548 | -35.5587 | -35.9487
25.4609 | -35.5591 | -35.9542
25.467 | -35.5594 | -35.9593
25.4731 | -35.5598 | -35.9651
25.4792 | -35.5601 | -35.9706
25.4853 | -35.5605 | -35.9763
25.4914 | -35.5609 | -35.9821
25.4975 | -35.5612 | -35.9884
25.5036 | -35.5615 | -35.9943
25.5097 | -35.5618 | -35.9994
25.5158 | -35.5622 | -36.0047
25.5219 | -35.5626 | -36.0099
25.528 | -35.563 | -36.0149
25.5341 | -35.5634 | -36.0201
25.5402 | -35.5637 | -36.0258
25.5463 | -35.5641 | -36.0317
25.5524 | -35.5645 | -36.0376
25.5585 | -35.5649 | -36.0434
25.5646 | -35.5651 | -36.0491
25.5707 | -35.5656 | -36.0549
25.5769 | -35.566 | -36.0605
25.583 | -35.5663 | -36.0657
25.5891 | -35.5667 | -36.0714
25.5952 | -35.5671 | -36.0771
25.6013 | -35.5675 | -36.0828
25.6074 | -35.5678 | -36.0883
25.6135 | -35.5682 | -36.0943
25.6196 | -35.5686 | -36.1002
25.6257 | -35.569 | -36.1063
25.6318 | -35.5694 | -36.1116
25.6379 | -35.5697 | -36.1176
25.644 | -35.5701 | -36.1234
25.6501 | -35.5704 | -36.1293
25.6562 | -35.5707 | -36.1352
25.6623 | -35.5712 | -36.1406
25.6684 | -35.5715 | -36.1465
25.6745 | -35.572 | -36.152
25.6806 | -35.5723 | -36.1576
25.6867 | -35.5727 | -36.1633
25.6928 | -35.573 | -36.1688
25.6989 | -35.5734 | -36.1739
25.705 | -35.5738 | -36.1797
25.7111 | -35.5741 | -36.1847
25.7173 | -35.5745 | -36.1897
25.7234 | -35.575 | -36.195
25.7295 | -35.5753 | -36.2003
25.7356 | -35.5757 | -36.2056
25.7417 | -35.5761 | -36.2112
25.7478 | -35.5764 | -36.216
25.7539 | -35.5767 | -36.2206
25.76 | -35.577 | -36.2264
25.7661 | -35.5774 | -36.2318
25.7722 | -35.5777 | -36.2371
25.7783 | -35.578 | -36.2423
25.7844 | -35.5784 | -36.2477
25.7905 | -35.5788 | -36.2534
25.7966 | -35.5792 | -36.2592
25.8027 | -35.5795 | -36.2644
25.8088 | -35.5799 | -36.2697
25.8149 | -35.5802 | -36.2752
25.821 | -35.5806 | -36.2802
25.8271 | -35.581 | -36.2857
25.8332 | -35.5814 | -36.291
25.8393 | -35.5817 | -36.2965
25.8454 | -35.5822 | -36.302
25.8515 | -35.5826 | -36.3074
25.8576 | -35.583 | -36.3121
25.8638 | -35.5833 | -36.3176
25.8699 | -35.5837 | -36.3231
25.876 | -35.584 | -36.3287
25.8821 | -35.5844 | -36.3343
25.8882 | -35.5847 | -36.3397
25.8943 | -35.5852 | -36.3456
25.9004 | -35.5855 | -36.3508
25.9065 | -35.586 | -36.3567
25.9126 | -35.5863 | -36.3624
25.9187 | -35.5867 | -36.3682
25.9248 | -35.5871 | -36.3735
25.9309 | -35.5874 | -36.3788
25.937 | -35.5879 | -36.3842
25.9431 | -35.5882 | -36.3896
25.9492 | -35.5885 | -36.3946
25.9553 | -35.5889 | -36.4001
25.9614 | -35.5892 | -36.4054
25.9675 | -35.5896 | -36.4106
25.9736 | -35.59 | -36.4161
25.9797 | -35.5903 | -36.4215
25.9858 | -35.5907 | -36.4275
25.9919 | -35.5911 | -36.4333
25.998 | -35.5914 | -36.4393
26.0042 | -35.5918 | -36.445
26.0103 | -35.5921 | -36.4507
26.0164 | -35.5925 | -36.4557
26.0225 | -35.593 | -36.4611
26.0286 | -35.5933 | -36.4661
26.0347 | -35.5937 | -36.4707
26.0408 | -35.5941 | -36.4762
26.0469 | -35.5945 | -36.4817
26.053 | -35.5948 | -36.4872
26.0591 | -35.5952 | -36.4924
26.0652 | -35.5955 | -36.4969
26.0713 | -35.5959 | -36.502
26.0774 | -35.5962 | -36.5073
26.0835 | -35.5966 | -36.5126
26.0896 | -35.597 | -36.5177
26.0957 | -35.5973 | -36.5228
26.1018 | -35.5978 | -36.5284
26.1079 | -35.5982 | -36.5337
26.114 | -35.5986 | -36.5386
26.1201 | -35.5989 | -36.5438
26.1262 | -35.5993 | -36.5494
26.1323 | -35.5996 | -36.5551
26.1384 | -35.5999 | -36.56
26.1445 | -35.6003 | -36.5657
26.1507 | -35.6006 | -36.5716
26.1568 | -35.6011 | -36.5769
26.1629 | -35.6015 | -36.582
26.169 | -35.6018 | -36.5875
26.1751 | -35.6021 | -36.5918
26.1812 | -35.6025 | -36.5974
26.1873 | -35.6028 | -36.6028
26.1934 | -35.6032 | -36.6082
26.1995 | -35.6036 | -36.6141
26.2056 | -35.604 | -36.6196
26.2117 | -35.6043 | -36.6249
26.2178 | -35.6046 | -36.6306
26.2239 | -35.605 | -36.6361
26.23 | -35.6053 | -36.6416
26.2361 | -35.6056 | -36.6472
26.2422 | -35.606 | -36.6517
26.2483 | -35.6064 | -36.657
26.2544 | -35.6067 | -36.662
26.2605 | -35.607 | -36.6679
26.2666 | -35.6074 | -36.6736
26.2727 | -35.6078 | -36.6788
26.2788 | -35.6081 | -36.6847
26.2849 | -35.6084 | -36.6898
26.2911 | -35.6088 | -36.6951
26.2972 | -35.6092 | -36.7004
26.3033 | -35.6095 | -36.7056
26.3094 | -35.6099 | -36.7108
26.3155 | -35.6103 | -36.7165
26.3216 | -35.6107 | -36.7219
26.3277 | -35.611 | -36.7277
26.3338 | -35.6114 | -36.7332
26.3399 | -35.6118 | -36.7385
26.346 | -35.6123 | -36.7431
26.3521 | -35.6127 | -36.7487
26.3582 | -35.613 | -36.7539
26.3643 | -35.6134 | -36.7593
26.3704 | -35.6138 | -36.7646
26.3765 | -35.6143 | -36.77
26.3826 | -35.6147 | -36.7745
26.3887 | -35.6151 | -36.7801
26.3948 | -35.6154 | -36.7848
26.4009 | -35.6158 | -36.7902
26.407 | -35.6161 | -36.7949
26.4131 | -35.6166 | -36.8001
26.4192 | -35.617 | -36.8051
26.4253 | -35.6173 | -36.8106
26.4314 | -35.6176 | -36.8158
26.4376 | -35.618 | -36.8213
26.4437 | -35.6183 | -36.8263
26.4498 | -35.6187 | -36.8318
26.4559 | -35.619 | -36.837
26.462 | -35.6193 | -36.842
26.4681 | -35.6197 | -36.8476
26.4742 | -35.62 | -36.8535
26.4803 | -35.6204 | -36.8591
26.4864 | -35.6207 | -36.8643
26.4925 | -35.621 | -36.8698
26.4986 | -35.6213 | -36.8752
26.5047 | -35.6217 | -36.8798
26.5108 | -35.6221 | -36.8852
26.5169 | -35.6224 | -36.8908
26.523 | -35.6227 | -36.8955
26.5291 | -35.6231 | -36.9009
26.5352 | -35.6234 | -36.9058
26.5413 | -35.6237 | -36.9115
26.5474 | -35.6241 | -36.9164
26.5535 | -35.6243 | -36.9218
26.5596 | -35.6246 | -36.9268
26.5657 | -35.6249 | -36.932
26.5718 | -35.6253 | -36.9376
26.578 | -35.6256 | -36.9428
26.5841 | -35.626 | -36.9476
26.5902 | -35.6263 | -36.9534
26.5963 | -35.6265 | -36.9584
26.6024 | -35.627 | -36.9635
26.6085 | -35.6273 | -36.9688
26.6146 | -35.6276 | -36.9734
26.6207 | -35.628 | -36.9785
26.6268 | -35.6282 | -36.9836
26.6329 | -35.6285 | -36.9887
26.639 | -35.6289 | -36.9943
26.6451 | -35.6293 | -36.9991
26.6512 | -35.6297 | -37.0046
26.6573 | -35.6299 | -37.0097
26.6634 | -35.6302 | -37.0152
26.6695 | -35.6305 | -37.0207
26.6756 | -35.6307 | -37.0262
26.6817 | -35.631 | -37.0317
26.6878 | -35.6313 | -37.0374
26.6939 | -35.6316 | -37.0425
26.7 | -35.632 | -37.047
26.7061 | -35.6323 | -37.0514
26.7122 | -35.6326 | -37.0567
26.7183 | -35.6328 | -37.0625
26.7245 | -35.6331 | -37.0687
26.7306 | -35.6334 | -37.0742
26.7367 | -35.6337 | -37.0791
26.7428 | -35.634 | -37.0845
26.7489 | -35.6344 | -37.0898
26.755 | -35.6348 | -37.0955
26.7611 | -35.6351 | -37.101
26.7672 | -35.6354 | -37.1063
26.7733 | -35.6359 | -37.1117
26.7794 | -35.6362 | -37.1178
26.7855 | -35.6365 | -37.1231
26.7916 | -35.6368 | -37.1282
26.7977 | -35.6372 | -37.1331
26.8038 | -35.6375 | -37.1381
26.8099 | -35.6379 | -37.1429
26.816 | -35.6382 | -37.1483
26.8221 | -35.6385 | -37.1531
26.8282 | -35.6389 | -37.1579
26.8343 | -35.6393 | -37.1635
26.8404 | -35.6396 | -37.1694
26.8465 | -35.6401 | -37.1745
26.8526 | -35.6404 | -37.18
26.8587 | -35.6407 | -37.1855
26.8649 | -35.6411 | -37.1905
26.871 | -35.6414 | -37.1956
26.8771 | -35.6417 | -37.2001
26.8832 | -35.642 | -37.2059
26.8893 | -35.6423 | -37.2111
26.8954 | -35.6426 | -37.2172
26.9015 | -35.6429 | -37.2229
26.9076 | -35.6434 | -37.2274
26.9137 | -35.6437 | -37.2322
26.9198 | -35.644 | -37.2378
26.9259 | -35.6445 | -37.2432
26.932 | -35.6448 | -37.2484
26.9381 | -35.645 | -37.2537
26.9442 | -35.6454 | -37.2588
26.9503 | -35.6458 | -37.2637
26.9564 | -35.6461 | -37.2693
26.9625 | -35.6464 | -37.2746
26.9686 | -35.6468 | -37.28
26.9747 | -35.647 | -37.2854
26.9808 | -35.6474 | -37.2904
26.9869 | -35.6477 | -37.2962
26.993 | -35.6481 | -37.3014
26.9991 | -35.6484 | -37.3062
27.0052 | -35.6486 | -37.3119
27.0114 | -35.649 | -37.3173
27.0175 | -35.6494 | -37.3231
27.0236 | -35.6497 | -37.3282
27.0297 | -35.65 | -37.3337
27.0358 | -35.6504 | -37.3391
27.0419 | -35.6508 | -37.3443
27.048 | -35.6512 | -37.3499
27.0541 | -35.6516 | -37.354
27.0602 | -35.6518 | -37.3589
27.0663 | -35.6522 | -37.3638
27.0724 | -35.6526 | -37.3689
27.0785 | -35.653 | -37.3736
27.0846 | -35.6533 | -37.3783
27.0907 | -35.6537 | -37.383
27.0968 | -35.6541 | -37.3885
27.1029 | -35.6544 | -37.3941
27.109 | -35.6548 | -37.3996
27.1151 | -35.655 | -37.4046
27.1212 | -35.6554 | -37.4099
27.1273 | -35.6558 | -37.4146
27.1334 | -35.6562 | -37.4193
27.1395 | -35.6564 | -37.425
27.1456 | -35.6567 | -37.4303
27.1518 | -35.6571 | -37.4353
27.1579 | -35.6574 | -37.441
27.164 | -35.6578 | -37.4462
27.1701 | -35.6582 | -37.451
27.1762 | -35.6585 | -37.4562
27.1823 | -35.6588 | -37.4618
27.1884 | -35.659 | -37.4672
27.1945 | -35.6593 | -37.4726
27.2006 | -35.6596 | -37.4777
27.2067 | -35.66 | -37.4834
27.2128 | -35.6603 | -37.4887
27.2189 | -35.6606 | -37.494
27.225 | -35.661 | -37.4997
27.2311 | -35.6615 | -37.5042
27.2372 | -35.6618 | -37.5086
27.2433 | -35.6622 | -37.5141
27.2494 | -35.6625 | -37.5193
27.2555 | -35.6629 | -37.5241
27.2616 | -35.6633 | -37.5286
27.2677 | -35.6637 | -37.5335
27.2738 | -35.664 | -37.5386
27.2799 | -35.6644 | -37.5435
27.286 | -35.6648 | -37.5478
27.2921 | -35.6651 | -37.5529
27.2983 | -35.6655 | -37.5575
27.3044 | -35.6659 | -37.5627
27.3105 | -35.6663 | -37.5673
27.3166 | -35.6666 | -37.5716
27.3227 | -35.667 | -37.5765
27.3288 | -35.6673 | -37.5818
27.3349 | -35.6677 | -37.587
27.341 | -35.668 | -37.5923
27.3471 | -35.6683 | -37.5971
27.3532 | -35.6686 | -37.6019
27.3593 | -35.669 | -37.6078
27.3654 | -35.6694 | -37.6131
27.3715 | -35.6697 | -37.6173
27.3776 | -35.6701 | -37.6225
27.3837 | -35.6703 | -37.6277
27.3898 | -35.6706 | -37.6326
27.3959 | -35.6709 | -37.6381
27.402 | -35.6712 | -37.643
27.4081 | -35.6716 | -37.6484
27.4142 | -35.672 | -37.6541
27.4203 | -35.6723 | -37.6594
27.4264 | -35.6727 | -37.6645
27.4325 | -35.673 | -37.6694
27.4387 | -35.6734 | -37.6748
27.4448 | -35.6737 | -37.6804
27.4509 | -35.6741 | -37.6858
27.457 | -35.6744 | -37.691
27.4631 | -35.6748 | -37.6961
27.4692 | -35.675 | -37.7008
27.4753 | -35.6753 | -37.7055
27.4814 | -35.6756 | -37.7104
27.4875 | -35.676 | -37.716
27.4936 | -35.6764 | -37.7209
27.4997 | -35.6766 | -37.7262
27.5058 | -35.6769 | -37.7314
27.5119 | -35.6773 | -37.7363
27.518 | -35.6775 | -37.7416
27.5241 | -35.6779 | -37.7463
27.5302 | -35.6782 | -37.7513
27.5363 | -35.6785 | -37.7565
27.5424 | -35.6788 | -37.7616
27.5485 | -35.6791 | -37.7664
27.5546 | -35.6795 | -37.7713
27.5607 | -35.6799 | -37.7766
27.5668 | -35.6802 | -37.7816
27.5729 | -35.6805 | -37.7861
27.5791 | -35.6809 | -37.7911
27.5852 | -35.6812 | -37.7961
27.5913 | -35.6814 | -37.8013
27.5974 | -35.6818 | -37.8061
27.6035 | -35.6822 | -37.8112
27.6096 | -35.6825 | -37.8167
27.6157 | -35.6828 | -37.8225
27.6218 | -35.6831 | -37.8273
27.6279 | -35.6834 | -37.8324
27.634 | -35.6838 | -37.837
27.6401 | -35.684 | -37.8422
27.6462 | -35.6844 | -37.8472
27.6523 | -35.6847 | -37.8519
27.6584 | -35.6851 | -37.8568
27.6645 | -35.6854 | -37.8624
27.6706 | -35.6857 | -37.8675
27.6767 | -35.6861 | -37.8724
27.6828 | -35.6864 | -37.8773
27.6889 | -35.6866 | -37.8822
27.695 | -35.6869 | -37.8874
27.7011 | -35.6872 | -37.8921
27.7072 | -35.6875 | -37.8973
27.7133 | -35.6878 | -37.9025
27.7194 | -35.6882 | -37.9083
27.7256 | -35.6885 | -37.9132
27.7317 | -35.6887 | -37.9181
27.7378 | -35.689 | -37.9234
27.7439 | -35.6894 | -37.9283
27.75 | -35.6896 | -37.9331
27.7561 | -35.69 | -37.9381
27.7622 | -35.6902 | -37.9435
27.7683 | -35.6906 | -37.9488
27.7744 | -35.6909 | -37.954
27.7805 | -35.6912 | -37.9595
27.7866 | -35.6916 | -37.9643
27.7927 | -35.6919 | -37.9694
27.7988 | -35.6923 | -37.9745
27.8049 | -35.6926 | -37.9791
27.811 | -35.6929 | -37.9843
27.8171 | -35.6933 | -37.989
27.8232 | -35.6936 | -37.9938
27.8293 | -35.6938 | -37.999
27.8354 | -35.6942 | -38.0042
27.8415 | -35.6945 | -38.0092
27.8476 | -35.6948 | -38.0143
27.8537 | -35.6952 | -38.0197
27.8598 | -35.6956 | -38.0245
27.866 | -35.6959 | -38.0294
27.8721 | -35.6962 | -38.0343
27.8782 | -35.6965 | -38.0396
27.8843 | -35.6968 | -38.0452
27.8904 | -35.6971 | -38.0498
27.8965 | -35.6974 | -38.0549
27.9026 | -35.6978 | -38.0608
27.9087 | -35.6981 | -38.066
27.9148 | -35.6984 | -38.0709
27.9209 | -35.6988 | -38.0755
27.927 | -35.6991 | -38.0802
27.9331 | -35.6995 | -38.085
27.9392 | -35.6998 | -38.0896
27.9453 | -35.7001 | -38.0948
27.9514 | -35.7004 | -38.1002
27.9575 | -35.7008 | -38.1052
27.9636 | -35.7012 | -38.1109
27.9697 | -35.7015 | -38.116
27.9758 | -35.7019 | -38.1211
27.9819 | -35.7023 | -38.1266
27.988 | -35.7026 | -38.1316
27.9941 | -35.7029 | -38.1366
28.0002 | -35.7031 | -38.1414
28.0063 | -35.7034 | -38.1463
28.0125 | -35.7037 | -38.1513
28.0186 | -35.7041 | -38.1564
28.0247 | -35.7044 | -38.1617
28.0308 | -35.7048 | -38.1669
28.0369 | -35.7051 | -38.1719
28.043 | -35.7054 | -38.177
28.0491 | -35.7058 | -38.182
28.0552 | -35.7061 | -38.1873
28.0613 | -35.7064 | -38.1922
28.0674 | -35.7066 | -38.1976
28.0735 | -35.707 | -38.2025

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.3383 | -42.818
21.4908 | -35.5048 | -35.2532
16.4948 | -34.875 | -27.8238
12.6602 | -34.5097 | -21.5133
9.7171 | -34.2906 | -16.5633
7.4582 | -34.1426 | -12.5733
5.7244 | -34.0411 | -9.3252
4.3936 | -33.9696 | -6.7207
3.3722 | -33.9601 | -4.663
2.5883 | -33.9729 | -3.0984
1.9866 | -33.9677 | -1.8553
1.5248 | -33.9619 | -0.85351
1.1703 | -33.9736 | -0.20209
1 | -33.9776 | 0.12097
0.89822 | -33.9803 | 0.39641
0.68941 | -33.9662 | 0.97935
0.52919 | -33.9671 | 1.4325
0.40612 | -33.9688 | 2.0409
0.31171 | -33.9633 | 2.4969
0.23928 | -33.9793 | 3.2221
0.18358 | -33.9839 | 3.644
0.14091 | -33.9872 | 4.5354
0.10819 | -33.9947 | 5.7032
0.083017 | -34.0375 | 7.1835
0.063705 | -34.0873 | 9.0653
0.048923 | -34.1616 | 11.5706
0.037479 | -34.2712 | 14.7699
0.028801 | -34.4554 | 18.7982
0.022078 | -34.7764 | 24.1754
0.017023 | -35.2609 | 30.6055
0.013018 | -36.0778 | 38.3053
0.010014 | -37.1939 | 46.21

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003658 | -11.2942
0.007316 | -28.2889
0.010974 | -36.4859
0.014633 | -40.632
0.018291 | -43.7226
0.021949 | -46.0709
0.025607 | -47.3185
0.029265 | -47.59
0.032923 | -48.4002
0.036581 | -49.9334
0.04024 | -51.6628
0.043898 | -53.8164
0.047556 | -56.2866
0.051214 | -56.9468
0.054872 | -56.6228
0.05853 | -57.0615
0.062188 | -58.0996
0.065847 | -58.7746
0.069505 | -58.637
0.073163 | -60.3472
0.076821 | -60.4431
0.080479 | -60.0834
0.084137 | -60.544
0.087796 | -60.9919
0.091454 | -62.1712
0.095112 | -63.6916
0.09877 | -64.477
0.10243 | -65.107
0.10609 | -63.8633
0.10974 | -63.8711
0.1134 | -65.165
0.11706 | -66.0873
0.12072 | -65.794
0.12438 | -66.4102
0.12804 | -67.5343
0.13169 | -67.9728
0.13535 | -68.7405
0.13901 | -68.2984
0.14267 | -68.1627
0.14633 | -67.5026
0.14998 | -68.4527
0.15364 | -68.5027
0.1573 | -69.0043
0.16096 | -70.6462
0.16462 | -70.8785
0.16828 | -70.3361
0.17193 | -69.9242
0.17559 | -70.5329
0.17925 | -69.9901
0.18291 | -69.925
0.18657 | -70.4865
0.19022 | -70.8606
0.19388 | -71.1145
0.19754 | -72.2577
0.2012 | -71.3328
0.20486 | -71.6169
0.20851 | -71.7388
0.21217 | -71.2369
0.21583 | -71.4826
0.21949 | -72.477
0.22315 | -72.5661
0.22681 | -72.084
0.23046 | -72.4082
0.23412 | -73.4697
0.23778 | -73.5705
0.24144 | -73.2201
0.2451 | -73.2689
0.24875 | -73.0934
0.25241 | -73.8032
0.25607 | -73.425
0.25973 | -74.1848
0.26339 | -74.2664
0.26704 | -74.4049
0.2707 | -75.4206
0.27436 | -74.7878
0.27802 | -74.5758
0.28168 | -75.9037
0.28534 | -76.3155
0.28899 | -74.5888
0.29265 | -75.1216
0.29631 | -75.9995
0.29997 | -75.956
0.30363 | -75.9072
0.30728 | -76.012
0.31094 | -75.3564
0.3146 | -76.2439
0.31826 | -76.6619
0.32192 | -76.331
0.32558 | -76.5272
0.32923 | -77.4727
0.33289 | -76.2264
0.33655 | -75.7687
0.34021 | -76.8507
0.34387 | -77.0606
0.34752 | -76.0121
0.35118 | -77.8342
0.35484 | -76.4547
0.3585 | -76.3715
0.36216 | -76.1672
0.36582 | -76.5885
0.36947 | -78.031
0.37313 | -78.8191
0.37679 | -79.289
0.38045 | -78.0651
0.3841 | -77.9149
0.38776 | -77.9308
0.39142 | -79.8857
0.39508 | -78.8073
0.39874 | -78.3787
0.4024 | -79.109
0.40605 | -78.8658
0.40971 | -77.7889
0.41337 | -78.5388
0.41703 | -79.3301
0.42069 | -79.2233
0.42435 | -79.2556
0.428 | -79.4351
0.43166 | -80.0764
0.43532 | -80.6107
0.43898 | -79.9443
0.44264 | -79.7971
0.44629 | -79.2686
0.44995 | -78.8465
0.45361 | -79.3401
0.45727 | -79.8127
0.46093 | -79.8805
0.46459 | -80.1865
0.46824 | -79.7277
0.4719 | -79.9212
0.47556 | -81.6312
0.47922 | -80.3025
0.48287 | -79.9396
0.48653 | -80.0679
0.49019 | -80.271
0.49385 | -81.4907
0.49751 | -80.8755
0.50117 | -80.8636
0.50482 | -80.6835
0.50848 | -81.127
0.51214 | -80.6123
0.5158 | -80.8299
0.51946 | -80.3876
0.52311 | -80.3465
0.52677 | -81.157
0.53043 | -81.6839
0.53409 | -81.1561
0.53775 | -80.4104
0.54141 | -80.5568
0.54506 | -81.6833
0.54872 | -82.3076
0.55238 | -81.6179
0.55604 | -81.1088
0.5597 | -81.4511
0.56336 | -81.9067
0.56701 | -81.8052
0.57067 | -81.6166
0.57433 | -81.4263
0.57799 | -81.229
0.58164 | -81.2138
0.5853 | -81.6287
0.58896 | -82.2522
0.59262 | -82.3833
0.59628 | -82.16
0.59994 | -82.3611
0.60359 | -82.8145
0.60725 | -82.6596
0.61091 | -82.2284
0.61457 | -82.2526
0.61823 | -82.6096
0.62189 | -82.7254
0.62554 | -82.7298
0.6292 | -83.1731
0.63286 | -83.627
0.63652 | -83.4233
0.64018 | -83.1153
0.64383 | -83.4802
0.64749 | -84.1626
0.65115 | -84.0351
0.65481 | -83.5548
0.65847 | -83.4542
0.66212 | -83.3721
0.66578 | -83.1583
0.66944 | -83.0217
0.6731 | -83.1187
0.67676 | -82.9625
0.68041 | -82.591
0.68407 | -82.6251
0.68773 | -82.9761
0.69139 | -83.216
0.69505 | -83.5901
0.69871 | -83.9426
0.70236 | -83.6998
0.70602 | -82.9363
0.70968 | -82.718
0.71334 | -83.3085
0.717 | -83.9424
0.72066 | -84.0237
0.72431 | -84.5391
0.72797 | -84.9753
0.73163 | -84.6521
0.73529 | -84.5876
0.73895 | -84.6001
0.7426 | -84.6219
0.74626 | -84.952
0.74992 | -85.065
0.75358 | -84.4893
0.75724 | -83.8703
0.76089 | -83.6103
0.76455 | -83.611
0.76821 | -83.8249
0.77187 | -84.2745
0.77553 | -84.6359
0.77919 | -84.6969
0.78284 | -84.5462
0.7865 | -84.9095
0.79016 | -85.6247
0.79382 | -85.8856
0.79748 | -85.4858
0.80113 | -85.2713
0.80479 | -85.4963
0.80845 | -85.3166
0.81211 | -84.9812
0.81577 | -85.3649
0.81942 | -85.9043
0.82308 | -86.4734
0.82674 | -86.6538
0.8304 | -85.8027
0.83406 | -84.9653
0.83771 | -84.952
0.84137 | -85.1581
0.84503 | -85.0417
0.84869 | -85.1735
0.85235 | -85.95
0.85601 | -86.5803
0.85966 | -86.3473
0.86332 | -86.0628
0.86698 | -85.9569
0.87064 | -85.9501
0.8743 | -86.0542
0.87796 | -86.4809
0.88161 | -86.7499
0.88527 | -86.1344
0.88893 | -85.5497
0.89259 | -86.0261
0.89625 | -86.4484
0.8999 | -86.093
0.90356 | -85.8571
0.90722 | -85.9192
0.91088 | -86.1232
0.91454 | -86.1005
0.91819 | -86.1291
0.92185 | -86.1764
0.92551 | -85.9877
0.92917 | -86.0105
0.93283 | -86.2539
0.93649 | -86.2584
0.94014 | -86.1056
0.9438 | -86.1875
0.94746 | -87.0331
0.95112 | -87.8227
0.95478 | -87.7795
0.95843 | -87.356
0.96209 | -86.6693
0.96575 | -86.4264
0.96941 | -86.5295
0.97307 | -86.6928
0.97672 | -86.8312
0.98038 | -86.8087
0.98404 | -86.8976
0.9877 | -86.8754
0.99136 | -86.58
0.99502 | -86.2135
0.99867 | -86.3876
1.0023 | -86.9883
1.006 | -87.4313
1.0096 | -87.6304
1.0133 | -87.3477
1.017 | -87.1106
1.0206 | -87.5694
1.0243 | -87.8148
1.0279 | -87.3832
1.0316 | -87.07
1.0353 | -87.4189
1.0389 | -88.0439
1.0426 | -88.1998
1.0462 | -87.9985
1.0499 | -87.7582
1.0535 | -87.4499
1.0572 | -87.3447
1.0609 | -87.4725
1.0645 | -87.7633
1.0682 | -88.0689
1.0718 | -88.2464
1.0755 | -88.3191
1.0792 | -88.4258
1.0828 | -88.5295
1.0865 | -88.4722
1.0901 | -88.3886
1.0938 | -88.4461
1.0974 | -88.5291
1.1011 | -88.5422
1.1048 | -88.5242
1.1084 | -88.4648
1.1121 | -88.4168
1.1157 | -88.5804
1.1194 | -89.0255
1.1231 | -89.4219
1.1267 | -89.4149
1.1304 | -89.0818
1.134 | -88.6485
1.1377 | -88.3247
1.1413 | -88.2499
1.145 | -88.3122
1.1487 | -88.3895
1.1523 | -88.524
1.156 | -88.6977
1.1596 | -88.7979
1.1633 | -88.8295
1.1669 | -88.7602
1.1706 | -88.5344
1.1743 | -88.2959
1.1779 | -88.1104
1.1816 | -87.9312
1.1852 | -87.911
1.1889 | -88.1925
1.1926 | -88.6912
1.1962 | -89.1547
1.1999 | -89.2922
1.2035 | -89.1613
1.2072 | -89.1287
1.2108 | -89.2618
1.2145 | -89.3022
1.2182 | -89.1362
1.2218 | -88.8912
1.2255 | -88.698
1.2291 | -88.7021
1.2328 | -88.9804
1.2365 | -89.3366
1.2401 | -89.5424
1.2438 | -89.5767
1.2474 | -89.4868
1.2511 | -89.3546
1.2547 | -89.2533
1.2584 | -89.2573
1.2621 | -89.3314
1.2657 | -89.4025
1.2694 | -89.5336
1.273 | -89.6785
1.2767 | -89.6106
1.2804 | -89.2703
1.284 | -88.915
1.2877 | -88.881
1.2913 | -89.2869
1.295 | -89.8735
1.2986 | -90.0811
1.3023 | -89.6677
1.306 | -89.0701
1.3096 | -88.6975
1.3133 | -88.6726
1.3169 | -89.0139
1.3206 | -89.5302
1.3242 | -89.7349
1.3279 | -89.4664
1.3316 | -89.1601
1.3352 | -89.1628
1.3389 | -89.4902
1.3425 | -90.0482
1.3462 | -90.5402
1.3499 | -90.3843
1.3535 | -89.727
1.3572 | -89.3319
1.3608 | -89.3632
1.3645 | -89.5194
1.3681 | -89.6828
1.3718 | -89.9313
1.3755 | -90.2232
1.3791 | -90.4346
1.3828 | -90.5377
1.3864 | -90.5927
1.3901 | -90.5666
1.3938 | -90.4603
1.3974 | -90.5102
1.4011 | -90.8155
1.4047 | -91.0664
1.4084 | -90.9978
1.412 | -90.6853
1.4157 | -90.2876
1.4194 | -89.9518
1.423 | -89.8056
1.4267 | -89.9568
1.4303 | -90.3956
1.434 | -91.0011
1.4377 | -91.6287
1.4413 | -92.0418
1.445 | -92.0607
1.4486 | -91.8576
1.4523 | -91.5365
1.4559 | -91.0716
1.4596 | -90.6477
1.4633 | -90.4108
1.4669 | -90.336
1.4706 | -90.3754
1.4742 | -90.4795
1.4779 | -90.5995
1.4815 | -90.6446
1.4852 | -90.5676
1.4889 | -90.3845
1.4925 | -90.1685
1.4962 | -90.0919
1.4998 | -90.2474
1.5035 | -90.5554
1.5072 | -90.8303
1.5108 | -90.853
1.5145 | -90.6206
1.5181 | -90.3065
1.5218 | -89.9499
1.5254 | -89.6381
1.5291 | -89.6645
1.5328 | -90.1564
1.5364 | -90.8482
1.5401 | -91.234
1.5437 | -91.1045
1.5474 | -90.7467
1.5511 | -90.4463
1.5547 | -90.2734
1.5584 | -90.2602
1.562 | -90.4589
1.5657 | -90.8069
1.5693 | -91.1253
1.573 | -91.2048
1.5767 | -90.9886
1.5803 | -90.9673
1.584 | -91.0984
1.5876 | -91.3862
1.5913 | -91.6443
1.595 | -91.7413
1.5986 | -91.7269
1.6023 | -91.7178
1.6059 | -91.7253
1.6096 | -91.6486
1.6132 | -91.4457
1.6169 | -91.2003
1.6206 | -91.061
1.6242 | -91.1203
1.6279 | -91.3596
1.6315 | -91.6748
1.6352 | -91.893
1.6388 | -91.9254
1.6425 | -91.8566
1.6462 | -91.8502
1.6498 | -91.9639
1.6535 | -92.1033
1.6571 | -92.1447
1.6608 | -92.0911
1.6645 | -92.0418
1.6681 | -92.0319
1.6718 | -92.0239
1.6754 | -92.0405
1.6791 | -92.0889
1.6827 | -92.1333
1.6864 | -92.1431
1.6901 | -92.1407
1.6937 | -92.1424
1.6974 | -92.1599
1.701 | -92.2293
1.7047 | -92.3609
1.7084 | -92.4996
1.712 | -92.5674
1.7157 | -92.5506
1.7193 | -92.4466
1.723 | -92.2602
1.7266 | -92.0456
1.7303 | -91.8876
1.734 | -91.8211
1.7376 | -91.7791
1.7413 | -91.7088
1.7449 | -91.618
1.7486 | -91.5754
1.7523 | -91.6893
1.7559 | -91.9843
1.7596 | -92.3822
1.7632 | -92.7311
1.7669 | -92.9067
1.7705 | -92.9277
1.7742 | -92.8357
1.7779 | -92.7038
1.7815 | -92.5918
1.7852 | -92.5639
1.7888 | -92.6416
1.7925 | -92.7973
1.7961 | -92.9979
1.7998 | -93.2001
1.8035 | -93.3319
1.8071 | -93.3939
1.8108 | -93.4012
1.8144 | -93.3282
1.8181 | -93.1354
1.8218 | -92.8094
1.8254 | -92.4084
1.8291 | -92.0398
1.8327 | -91.771
1.8364 | -91.6337
1.84 | -91.6719
1.8437 | -91.9078
1.8474 | -92.2654
1.851 | -92.6199
1.8547 | -92.914
1.8583 | -93.1496
1.862 | -93.305
1.8657 | -93.3672
1.8693 | -93.2999
1.873 | -93.0825
1.8766 | -92.8187
1.8803 | -92.7021
1.8839 | -92.7865
1.8876 | -92.9681
1.8913 | -93.099
1.8949 | -93.1485
1.8986 | -93.1693
1.9022 | -93.1854
1.9059 | -93.1765
1.9096 | -93.1384
1.9132 | -93.0676
1.9169 | -92.9605
1.9205 | -92.8462
1.9242 | -92.8014
1.9278 | -92.8369
1.9315 | -92.8882
1.9352 | -92.9236
1.9388 | -92.9509
1.9425 | -92.9657
1.9461 | -92.9999
1.9498 | -93.1115
1.9534 | -93.3191
1.9571 | -93.5544
1.9608 | -93.7313
1.9644 | -93.7873
1.9681 | -93.7164
1.9717 | -93.5902
1.9754 | -93.5015
1.9791 | -93.5069
1.9827 | -93.6283
1.9864 | -93.7905
1.99 | -93.8617
1.9937 | -93.7617
1.9973 | -93.5492
2.001 | -93.3393
2.0047 | -93.252
2.0083 | -93.3784
2.012 | -93.7032
2.0156 | -94.0544
2.0193 | -94.2423
2.023 | -94.2358
2.0266 | -94.1501
2.0303 | -94.0635
2.0339 | -93.9269
2.0376 | -93.6494
2.0412 | -93.2813
2.0449 | -92.9601
2.0486 | -92.7513
2.0522 | -92.652
2.0559 | -92.6865
2.0595 | -92.902
2.0632 | -93.2687
2.0669 | -93.6319
2.0705 | -93.8397
2.0742 | -93.8723
2.0778 | -93.8265
2.0815 | -93.8173
2.0851 | -93.8863
2.0888 | -94.0325
2.0925 | -94.1578
2.0961 | -94.1602
2.0998 | -94.0714
2.1034 | -94.022
2.1071 | -94.1328
2.1107 | -94.3368
2.1144 | -94.4342
2.1181 | -94.4309
2.1217 | -94.3131
2.1254 | -94.1148
2.129 | -93.9074
2.1327 | -93.7685
2.1364 | -93.7465
2.14 | -93.847
2.1437 | -94.0229
2.1473 | -94.219
2.151 | -94.3726
2.1546 | -94.416
2.1583 | -94.2903
2.162 | -94.0457
2.1656 | -93.7854
2.1693 | -93.5998
2.1729 | -93.5077
2.1766 | -93.5082
2.1803 | -93.5922
2.1839 | -93.7542
2.1876 | -93.9593
2.1912 | -94.1381
2.1949 | -94.2349
2.1985 | -94.2247
2.2022 | -94.1028
2.2059 | -93.9109
2.2095 | -93.7087
2.2132 | -93.56
2.2168 | -93.4852
2.2205 | -93.4721
2.2242 | -93.491
2.2278 | -93.5284
2.2315 | -93.5844
2.2351 | -93.6697
2.2388 | -93.7852
2.2424 | -93.9078
2.2461 | -93.9804
2.2498 | -94.0026
2.2534 | -94.0086
2.2571 | -94.0433
2.2607 | -94.1268
2.2644 | -94.2369
2.2681 | -94.3444
2.2717 | -94.4091
2.2754 | -94.4092
2.279 | -94.3609
2.2827 | -94.3398
2.2863 | -94.4019
2.29 | -94.5529
2.2937 | -94.7613
2.2973 | -94.9613
2.301 | -95.0789
2.3046 | -95.0903
2.3083 | -95.0035
2.3119 | -94.8923
2.3156 | -94.8188
2.3193 | -94.7962
2.3229 | -94.7969
2.3266 | -94.7863
2.3302 | -94.7345
2.3339 | -94.6343
2.3376 | -94.4988
2.3412 | -94.3619
2.3449 | -94.2648
2.3485 | -94.2445
2.3522 | -94.3227
2.3558 | -94.4918
2.3595 | -94.7375
2.3632 | -95.0041
2.3668 | -95.2237
2.3705 | -95.3123
2.3741 | -95.2411
2.3778 | -95.0299
2.3815 | -94.7448
2.3851 | -94.4749
2.3888 | -94.2994
2.3924 | -94.2625
2.3961 | -94.3629
2.3997 | -94.5299
2.4034 | -94.6869
2.4071 | -94.7785
2.4107 | -94.7641
2.4144 | -94.6089
2.418 | -94.3724
2.4217 | -94.1621
2.4254 | -94.0886
2.429 | -94.193
2.4327 | -94.4516
2.4363 | -94.7966
2.44 | -95.1422
2.4436 | -95.4015
2.4473 | -95.5068
2.451 | -95.4483
2.4546 | -95.2737
2.4583 | -95.0296
2.4619 | -94.7943
2.4656 | -94.6334
2.4692 | -94.5819
2.4729 | -94.6242
2.4766 | -94.73
2.4802 | -94.8543
2.4839 | -94.9384
2.4875 | -94.9597
2.4912 | -94.9214
2.4949 | -94.837
2.4985 | -94.7456
2.5022 | -94.6644
2.5058 | -94.6054
2.5095 | -94.5854
2.5131 | -94.6178
2.5168 | -94.6994
2.5205 | -94.817
2.5241 | -94.9495
2.5278 | -95.0145
2.5314 | -94.9578
2.5351 | -94.839
2.5388 | -94.78
2.5424 | -94.8662
2.5461 | -95.0927
2.5497 | -95.3735
2.5534 | -95.5896
2.557 | -95.6842
2.5607 | -95.687
2.5644 | -95.6468
2.568 | -95.6078
2.5717 | -95.5514
2.5753 | -95.4495
2.579 | -95.3288
2.5827 | -95.27
2.5863 | -95.3076
2.59 | -95.416
2.5936 | -95.5364
2.5973 | -95.5979
2.6009 | -95.5853
2.6046 | -95.5432
2.6083 | -95.5205
2.6119 | -95.5409
2.6156 | -95.5796
2.6192 | -95.5912
2.6229 | -95.5258
2.6265 | -95.3666
2.6302 | -95.1635
2.6339 | -94.9752
2.6375 | -94.9117
2.6412 | -94.8612
2.6448 | -94.8749
2.6485 | -94.9441
2.6522 | -95.066
2.6558 | -95.2305
2.6595 | -95.4041
2.6631 | -95.5486
2.6668 | -95.6392
2.6704 | -95.6896
2.6741 | -95.7016
2.6778 | -95.7006
2.6814 | -95.6921
2.6851 | -95.6762
2.6887 | -95.6523
2.6924 | -95.6217
2.6961 | -95.6029
2.6997 | -95.6017
2.7034 | -95.6226
2.707 | -95.6675
2.7107 | -95.7119
2.7143 | -95.7376
2.718 | -95.7355
2.7217 | -95.7275
2.7253 | -95.7281
2.729 | -95.7482
2.7326 | -95.7833
2.7363 | -95.832
2.74 | -95.8626
2.7436 | -95.8626
2.7473 | -95.8293
2.7509 | -95.7802
2.7546 | -95.7393
2.7582 | -95.7184
2.7619 | -95.7306
2.7656 | -95.783
2.7692 | -95.8697
2.7729 | -95.988
2.7765 | -96.1122
2.7802 | -96.2026
2.7838 | -96.2379
2.7875 | -96.2353
2.7912 | -96.2041
2.7948 | -96.1546
2.7985 | -96.093
2.8021 | -96.0311
2.8058 | -96.0067
2.8095 | -96.0304
2.8131 | -96.0935
2.8168 | -96.172
2.8204 | -96.2135
2.8241 | -96.191
2.8277 | -96.0937
2.8314 | -95.9395
2.8351 | -95.7771
2.8387 | -95.6497
2.8424 | -95.5662
2.846 | -95.5368
2.8497 | -95.5624
2.8534 | -95.6303
2.857 | -95.7263
2.8607 | -95.8256
2.8643 | -95.8937
2.868 | -95.9306
2.8716 | -95.9502
2.8753 | -95.9859
2.879 | -96.0528
2.8826 | -96.1473
2.8863 | -96.2523
2.8899 | -96.3576
2.8936 | -96.4501
2.8973 | -96.5155
2.9009 | -96.5544
2.9046 | -96.547
2.9082 | -96.4936
2.9119 | -96.4116
2.9155 | -96.3215
2.9192 | -96.2438
2.9229 | -96.1657
2.9265 | -96.0732
2.9302 | -95.9626
2.9338 | -95.8393
2.9375 | -95.7211
2.9411 | -95.6375
2.9448 | -95.6212
2.9485 | -95.6719
2.9521 | -95.7847
2.9558 | -95.9314
2.9594 | -96.0693
2.9631 | -96.1688
2.9668 | -96.2268
2.9704 | -96.2591
2.9741 | -96.2889
2.9777 | -96.3326
2.9814 | -96.4069
2.985 | -96.5189
2.9887 | -96.6691
2.9924 | -96.8571
2.996 | -97.0456
2.9997 | -97.1963
3.0033 | -97.2527
3.007 | -97.1912
3.0107 | -97.0053
3.0143 | -96.7654
3.018 | -96.5444
3.0216 | -96.3747
3.0253 | -96.2752
3.0289 | -96.2434
3.0326 | -96.2634
3.0363 | -96.3266
3.0399 | -96.4294
3.0436 | -96.5644
3.0472 | -96.6895
3.0509 | -96.7927
3.0546 | -96.8779
3.0582 | -96.9579
3.0619 | -97.0329
3.0655 | -97.1082
3.0692 | -97.1758
3.0728 | -97.2098
3.0765 | -97.1783
3.0802 | -97.0767
3.0838 | -96.9055
3.0875 | -96.7074
3.0911 | -96.5122
3.0948 | -96.3641
3.0985 | -96.2759
3.1021 | -96.2615
3.1058 | -96.3223
3.1094 | -96.4501
3.1131 | -96.6369
3.1167 | -96.8542
3.1204 | -97.0435
3.1241 | -97.1757
3.1277 | -97.2478
3.1314 | -97.2808
3.135 | -97.2843
3.1387 | -97.2883
3.1423 | -97.2965
3.146 | -97.3085
3.1497 | -97.2971
3.1533 | -97.2434
3.157 | -97.1419
3.1606 | -96.9981
3.1643 | -96.8399
3.168 | -96.708
3.1716 | -96.6074
3.1753 | -96.5477
3.1789 | -96.5379
3.1826 | -96.5868
3.1862 | -96.6787
3.1899 | -96.7878
3.1936 | -96.8877
3.1972 | -96.9662
3.2009 | -97.0156
3.2045 | -97.0493
3.2082 | -97.0736
3.2119 | -97.0988
3.2155 | -97.1291
3.2192 | -97.1761
3.2228 | -97.2261
3.2265 | -97.2572
3.2301 | -97.261
3.2338 | -97.2255
3.2375 | -97.1584
3.2411 | -97.0802
3.2448 | -97.0087
3.2484 | -96.9663
3.2521 | -96.9597
3.2557 | -96.9945
3.2594 | -97.0537
3.2631 | -97.1258
3.2667 | -97.1979
3.2704 | -97.2661
3.274 | -97.3149
3.2777 | -97.359
3.2814 | -97.3783
3.285 | -97.3507
3.2887 | -97.2872
3.2923 | -97.1951
3.296 | -97.0973
3.2996 | -97.0112
3.3033 | -96.9579
3.307 | -96.9491
3.3106 | -96.9813
3.3143 | -97.0402
3.3179 | -97.1158
3.3216 | -97.2086
3.3253 | -97.305
3.3289 | -97.4017
3.3326 | -97.4789
3.3362 | -97.548
3.3399 | -97.6088
3.3435 | -97.6606
3.3472 | -97.7217
3.3509 | -97.7987
3.3545 | -97.9084
3.3582 | -98.0275
3.3618 | -98.1414
3.3655 | -98.2356
3.3692 | -98.298
3.3728 | -98.3227
3.3765 | -98.3058
3.3801 | -98.2808
3.3838 | -98.2626
3.3874 | -98.2538
3.3911 | -98.2449
3.3948 | -98.2272
3.3984 | -98.1937
3.4021 | -98.1439
3.4057 | -98.0897
3.4094 | -98.0372
3.4131 | -97.9922
3.4167 | -97.9478
3.4204 | -97.8928
3.424 | -97.8334
3.4277 | -97.7866
3.4313 | -97.7508
3.435 | -97.7269
3.4387 | -97.7157
3.4423 | -97.7196
3.446 | -97.7296
3.4496 | -97.7425
3.4533 | -97.7622
3.4569 | -97.791
3.4606 | -97.8155
3.4643 | -97.8375
3.4679 | -97.8611
3.4716 | -97.8777
3.4752 | -97.8915
3.4789 | -97.9028
3.4826 | -97.9023
3.4862 | -97.8867
3.4899 | -97.851
3.4935 | -97.7931
3.4972 | -97.7238
3.5008 | -97.648
3.5045 | -97.5703
3.5082 | -97.4898
3.5118 | -97.4092
3.5155 | -97.3474
3.5191 | -97.3076
3.5228 | -97.2993
3.5265 | -97.3175
3.5301 | -97.3681
3.5338 | -97.4443
3.5374 | -97.5453
3.5411 | -97.6753
3.5447 | -97.7946
3.5484 | -97.8933
3.5521 | -97.9653
3.5557 | -98.0096
3.5594 | -98.0343
3.563 | -98.0531
3.5667 | -98.0737
3.5704 | -98.095
3.574 | -98.0978
3.5777 | -98.072
3.5813 | -98.018
3.585 | -97.9105
3.5886 | -97.795
3.5923 | -97.7069
3.596 | -97.6612
3.5996 | -97.683
3.6033 | -97.7634
3.6069 | -97.8759
3.6106 | -98.0228
3.6142 | -98.1922
3.6179 | -98.3732
3.6216 | -98.5505
3.6252 | -98.7004
3.6289 | -98.8174
3.6325 | -98.8811
3.6362 | -98.8933
3.6399 | -98.8577
3.6435 | -98.7678
3.6472 | -98.6332
3.6508 | -98.4809
3.6545 | -98.333
3.6581 | -98.2096
3.6618 | -98.123
3.6655 | -98.0745
3.6691 | -98.0723
3.6728 | -98.1037
3.6764 | -98.1488
3.6801 | -98.1972
3.6838 | -98.2378
3.6874 | -98.2515
3.6911 | -98.2335
3.6947 | -98.227
3.6984 | -98.2285
3.702 | -98.2343
3.7057 | -98.2523
3.7094 | -98.289
3.713 | -98.3267
3.7167 | -98.3626
3.7203 | -98.4094
3.724 | -98.4483
3.7277 | -98.4619
3.7313 | -98.4597
3.735 | -98.4433
3.7386 | -98.4034
3.7423 | -98.3573
3.7459 | -98.3227
3.7496 | -98.305
3.7533 | -98.3085
3.7569 | -98.3287
3.7606 | -98.364
3.7642 | -98.4078
3.7679 | -98.4493
3.7715 | -98.4739
3.7752 | -98.4875
3.7789 | -98.4871
3.7825 | -98.4689
3.7862 | -98.4425
3.7898 | -98.4103
3.7935 | -98.3632
3.7972 | -98.3037
3.8008 | -98.2422
3.8045 | -98.1828
3.8081 | -98.151
3.8118 | -98.1471
3.8154 | -98.1761
3.8191 | -98.24
3.8228 | -98.3207
3.8264 | -98.4073
3.8301 | -98.4833
3.8337 | -98.5341
3.8374 | -98.5646
3.8411 | -98.5846
3.8447 | -98.604
3.8484 | -98.6366
3.852 | -98.6917
3.8557 | -98.7755
3.8593 | -98.8788
3.863 | -98.9992
3.8667 | -99.1326
3.8703 | -99.2608
3.874 | -99.3775
3.8776 | -99.4761
3.8813 | -99.5446
3.885 | -99.5801
3.8886 | -99.5782
3.8923 | -99.5383
3.8959 | -99.4616
3.8996 | -99.3297
3.9032 | -99.1603
3.9069 | -98.9867
3.9106 | -98.8224
3.9142 | -98.6781
3.9179 | -98.5648
3.9215 | -98.4916
3.9252 | -98.4594
3.9288 | -98.4574
3.9325 | -98.4715
3.9362 | -98.4978
3.9398 | -98.5243
3.9435 | -98.5421
3.9471 | -98.5459
3.9508 | -98.5445
3.9545 | -98.5456
3.9581 | -98.56
3.9618 | -98.5993
3.9654 | -98.6661
3.9691 | -98.7565
3.9727 | -98.8669
3.9764 | -98.9881
3.9801 | -99.096
3.9837 | -99.1837
3.9874 | -99.2319
3.991 | -99.2559
3.9947 | -99.2607
3.9984 | -99.2481
4.002 | -99.2228
4.0057 | -99.1868
4.0093 | -99.1466
4.013 | -99.0976
4.0166 | -99.0419
4.0203 | -98.9783
4.024 | -98.9084
4.0276 | -98.8313
4.0313 | -98.7637
4.0349 | -98.7164
4.0386 | -98.6835
4.0423 | -98.6679
4.0459 | -98.6617
4.0496 | -98.6732
4.0532 | -98.6959
4.0569 | -98.7349
4.0605 | -98.7818
4.0642 | -98.8291
4.0679 | -98.8702
4.0715 | -98.894
4.0752 | -98.9022
4.0788 | -98.8951
4.0825 | -98.8924
4.0861 | -98.9047
4.0898 | -98.928
4.0935 | -98.9553
4.0971 | -98.9771
4.1008 | -98.9776
4.1044 | -98.9575
4.1081 | -98.9182
4.1118 | -98.8578
4.1154 | -98.7852
4.1191 | -98.7158
4.1227 | -98.6629
4.1264 | -98.6295
4.13 | -98.6172
4.1337 | -98.6301
4.1374 | -98.6596
4.141 | -98.7041
4.1447 | -98.751
4.1483 | -98.7979
4.152 | -98.8473
4.1557 | -98.9027
4.1593 | -98.9686
4.163 | -99.0556
4.1666 | -99.1537
4.1703 | -99.2532
4.1739 | -99.3481
4.1776 | -99.4235
4.1813 | -99.4593
4.1849 | -99.4535
4.1886 | -99.4247
4.1922 | -99.388
4.1959 | -99.3537
4.1996 | -99.3389
4.2032 | -99.3515
4.2069 | -99.3939
4.2105 | -99.4581
4.2142 | -99.5383
4.2178 | -99.6089
4.2215 | -99.5955
4.2252 | -99.5964
4.2288 | -99.5696
4.2325 | -99.5009
4.2361 | -99.4172
4.2398 | -99.3303
4.2435 | -99.2522
4.2471 | -99.2041
4.2508 | -99.1825
4.2544 | -99.1869
4.2581 | -99.2043
4.2617 | -99.2359
4.2654 | -99.2861
4.2691 | -99.3397
4.2727 | -99.388
4.2764 | -99.4291
4.28 | -99.4558
4.2837 | -99.4665
4.2873 | -99.463
4.291 | -99.4429
4.2947 | -99.4126
4.2983 | -99.381
4.302 | -99.3562
4.3056 | -99.327
4.3093 | -99.2975
4.313 | -99.2729
4.3166 | -99.2636
4.3203 | -99.2658
4.3239 | -99.2725
4.3276 | -99.2963
4.3312 | -99.3221
4.3349 | -99.3439
4.3386 | -99.3643
4.3422 | -99.3986
4.3459 | -99.441
4.3495 | -99.4885
4.3532 | -99.5448
4.3569 | -99.6007
4.3605 | -99.6568
4.3642 | -99.7069
4.3678 | -99.7429
4.3715 | -99.767
4.3751 | -99.7798
4.3788 | -99.7766
4.3825 | -99.758
4.3861 | -99.7167
4.3898 | -99.676
4.3934 | -99.6436
4.3971 | -99.6235
4.4008 | -99.6275
4.4044 | -99.6486
4.4081 | -99.6704
4.4117 | -99.6756
4.4154 | -99.6723
4.419 | -99.6566
4.4227 | -99.6179
4.4264 | -99.5702
4.43 | -99.5207
4.4337 | -99.4818
4.4373 | -99.4581
4.441 | -99.4593
4.4446 | -99.4814
4.4483 | -99.5254
4.452 | -99.5889
4.4556 | -99.6648
4.4593 | -99.7372
4.4629 | -99.7972
4.4666 | -99.8373
4.4703 | -99.8565
4.4739 | -99.8571
4.4776 | -99.8553
4.4812 | -99.8515
4.4849 | -99.837
4.4885 | -99.8228
4.4922 | -99.815
4.4959 | -99.7957
4.4995 | -99.7646
4.5032 | -99.7225
4.5068 | -99.6725
4.5105 | -99.6197
4.5142 | -99.5763
4.5178 | -99.5502
4.5215 | -99.5442
4.5251 | -99.5566
4.5288 | -99.5855
4.5324 | -99.63
4.5361 | -99.6731
4.5398 | -99.716
4.5434 | -99.7538
4.5471 | -99.7781
4.5507 | -99.7924
4.5544 | -99.8072
4.558 | -99.8245
4.5617 | -99.8481
4.5654 | -99.8846
4.569 | -99.9354
4.5727 | -99.9956
4.5763 | -100.064
4.58 | -100.1275
4.5837 | -100.1918
4.5873 | -100.2403
4.591 | -100.2562
4.5946 | -100.2368
4.5983 | -100.1863
4.6019 | -100.1027
4.6056 | -100.0036
4.6093 | -99.8986
4.6129 | -99.8028
4.6166 | -99.7313
4.6202 | -99.6821
4.6239 | -99.6606
4.6276 | -99.661
4.6312 | -99.6763
4.6349 | -99.6914
4.6385 | -99.7046
4.6422 | -99.7093
4.6458 | -99.7037
4.6495 | -99.6882
4.6532 | -99.6573
4.6568 | -99.624
4.6605 | -99.5967
4.6641 | -99.5859
4.6678 | -99.5882
4.6715 | -99.6134
4.6751 | -99.6641
4.6788 | -99.7248
4.6824 | -99.789
4.6861 | -99.8405
4.6897 | -99.8766
4.6934 | -99.888
4.6971 | -99.8856
4.7007 | -99.8836
4.7044 | -99.903
4.708 | -99.9384
4.7117 | -99.9993
4.7153 | -100.0839
4.719 | -100.1811
4.7227 | -100.2746
4.7263 | -100.356
4.73 | -100.4184
4.7336 | -100.4449
4.7373 | -100.4393
4.741 | -100.4021
4.7446 | -100.3483
4.7483 | -100.2663
4.7519 | -100.1908
4.7556 | -100.1113
4.7592 | -100.0346
4.7629 | -99.9574
4.7666 | -99.8865
4.7702 | -99.8275
4.7739 | -99.7759
4.7775 | -99.7356
4.7812 | -99.7146
4.7849 | -99.7122
4.7885 | -99.7214
4.7922 | -99.7476
4.7958 | -99.7902
4.7995 | -99.8425
4.8031 | -99.8907
4.8068 | -99.9461
4.8105 | -99.9988
4.8141 | -100.0423
4.8178 | -100.0791
4.8214 | -100.1139
4.8251 | -100.1501
4.8288 | -100.1894
4.8324 | -100.2285
4.8361 | -100.2679
4.8397 | -100.3109
4.8434 | -100.3521
4.847 | -100.3841
4.8507 | -100.4059
4.8544 | -100.4169
4.858 | -100.417
4.8617 | -100.4132
4.8653 | -100.4057
4.869 | -100.3987
4.8727 | -100.3987
4.8763 | -100.4007
4.88 | -100.4031
4.8836 | -100.3922
4.8873 | -100.3779
4.8909 | -100.3689
4.8946 | -100.3504
4.8983 | -100.337
4.9019 | -100.3369
4.9056 | -100.3511
4.9092 | -100.3725
4.9129 | -100.4072
4.9165 | -100.4458
4.9202 | -100.4853
4.9239 | -100.5294
4.9275 | -100.5869
4.9312 | -100.6482
4.9348 | -100.7145
4.9385 | -100.78
4.9422 | -100.8442
4.9458 | -100.8979
4.9495 | -100.9322
4.9531 | -100.9595
4.9568 | -100.9698
4.9604 | -100.9607
4.9641 | -100.9302
4.9678 | -100.8845
4.9714 | -100.8403
4.9751 | -100.7966
4.9787 | -100.7567
4.9824 | -100.7249
4.9861 | -100.7033
4.9897 | -100.693
4.9934 | -100.6953
4.997 | -100.6994
5.0007 | -100.699
5.0043 | -100.6981
5.008 | -100.7038
5.0117 | -100.713
5.0153 | -100.7391
5.019 | -100.7835
5.0226 | -100.8425
5.0263 | -100.9085
5.03 | -100.9725
5.0336 | -101.0322
5.0373 | -101.0663
5.0409 | -101.0805
5.0446 | -101.0711
5.0482 | -101.035
5.0519 | -100.9745
5.0556 | -100.8982
5.0592 | -100.8221
5.0629 | -100.7442
5.0665 | -100.6657
5.0702 | -100.6024
5.0738 | -100.5592
5.0775 | -100.5311
5.0812 | -100.5295
5.0848 | -100.5431
5.0885 | -100.5751
5.0921 | -100.6144
5.0958 | -100.6599
5.0995 | -100.7067
5.1031 | -100.7619
5.1068 | -100.8189
5.1104 | -100.8822
5.1141 | -100.9445
5.1177 | -100.9879
5.1214 | -101.0088
5.1251 | -101.0152
5.1287 | -101.0045
5.1324 | -100.9693
5.136 | -100.9016
5.1397 | -100.8233
5.1434 | -100.7354
5.147 | -100.644
5.1507 | -100.5717
5.1543 | -100.5257
5.158 | -100.4978
5.1616 | -100.5033
5.1653 | -100.5325
5.169 | -100.5824
5.1726 | -100.6468
5.1763 | -100.7242
5.1799 | -100.8069
5.1836 | -100.8877
5.1873 | -100.9601
5.1909 | -101.0328
5.1946 | -101.0879
5.1982 | -101.1238
5.2019 | -101.1429
5.2055 | -101.1471
5.2092 | -101.1459
5.2129 | -101.1535
5.2165 | -101.1658
5.2202 | -101.1855
5.2238 | -101.2077
5.2275 | -101.2198
5.2311 | -101.2212
5.2348 | -101.2098
5.2385 | -101.1778
5.2421 | -101.1303
5.2458 | -101.0698
5.2494 | -101.0083
5.2531 | -100.9502
5.2568 | -100.8961
5.2604 | -100.8598
5.2641 | -100.847
5.2677 | -100.8518
5.2714 | -100.8793
5.275 | -100.9351
5.2787 | -100.9783
5.2824 | -101.028
5.286 | -101.0734
5.2897 | -101.1196
5.2933 | -101.1576
5.297 | -101.1877
5.3007 | -101.2156
5.3043 | -101.2362
5.308 | -101.2481
5.3116 | -101.2534
5.3153 | -101.2493
5.3189 | -101.2393
5.3226 | -101.2231
5.3263 | -101.1993
5.3299 | -101.1695
5.3336 | -101.1331
5.3372 | -101.0988
5.3409 | -101.0563
5.3446 | -101.009
5.3482 | -100.9663
5.3519 | -100.9237
5.3555 | -100.8804
5.3592 | -100.8499
5.3628 | -100.821
5.3665 | -100.8027
5.3702 | -100.8006
5.3738 | -100.8093
5.3775 | -100.827
5.3811 | -100.8568
5.3848 | -100.8916
5.3884 | -100.9312
5.3921 | -100.9764
5.3958 | -101.0133
5.3994 | -101.042
5.4031 | -101.0615
5.4067 | -101.0645
5.4104 | -101.0533
5.4141 | -101.0302
5.4177 | -101.0064
5.4214 | -100.9763
5.425 | -100.9557
5.4287 | -100.9457
5.4323 | -100.949
5.436 | -100.9741
5.4397 | -101.0141
5.4433 | -101.063
5.447 | -101.1258
5.4506 | -101.1897
5.4543 | -101.2514
5.458 | -101.3137
5.4616 | -101.3732
5.4653 | -101.4181
5.4689 | -101.4479
5.4726 | -101.4572
5.4762 | -101.4431
5.4799 | -101.417
5.4836 | -101.3833
5.4872 | -101.3479
5.4909 | -101.3283
5.4945 | -101.3238
5.4982 | -101.3347
5.5019 | -101.3567
5.5055 | -101.3982
5.5092 | -101.4471
5.5128 | -101.4909
5.5165 | -101.5327
5.5201 | -101.5668
5.5238 | -101.59
5.5275 | -101.5931
5.5311 | -101.5721
5.5348 | -101.5358
5.5384 | -101.4876
5.5421 | -101.4208
5.5457 | -101.3472
5.5494 | -101.2831
5.5531 | -101.2284
5.5567 | -101.1983
5.5604 | -101.1871
5.564 | -101.1956
5.5677 | -101.2209
5.5714 | -101.2598
5.575 | -101.3061
5.5787 | -101.3594
5.5823 | -101.4129
5.586 | -101.4639
5.5896 | -101.5076
5.5933 | -101.5501
5.597 | -101.5899
5.6006 | -101.6118
5.6043 | -101.6144
5.6079 | -101.5895
5.6116 | -101.5425
5.6153 | -101.4739
5.6189 | -101.3755
5.6226 | -101.2624
5.6262 | -101.1396
5.6299 | -101.0182
5.6335 | -100.9011
5.6372 | -100.8041
5.6409 | -100.7245
5.6445 | -100.6756
5.6482 | -100.6573
5.6518 | -100.6707
5.6555 | -100.7135
5.6592 | -100.7784
5.6628 | -100.8584
5.6665 | -100.9504
5.6701 | -101.0477
5.6738 | -101.1474
5.6774 | -101.2373
5.6811 | -101.3195
5.6848 | -101.3939
5.6884 | -101.4533
5.6921 | -101.4954
5.6957 | -101.5147
5.6994 | -101.5148
5.7031 | -101.5038
5.7067 | -101.4788
5.7104 | -101.452
5.714 | -101.4226
5.7177 | -101.4032
5.7213 | -101.3943
5.725 | -101.3969
5.7287 | -101.409
5.7323 | -101.4294
5.736 | -101.4523
5.7396 | -101.4894
5.7433 | -101.5279
5.7469 | -101.5688
5.7506 | -101.6032
5.7543 | -101.6305
5.7579 | -101.6415
5.7616 | -101.6417
5.7652 | -101.6335
5.7689 | -101.6215
5.7726 | -101.6098
5.7762 | -101.6039
5.7799 | -101.5992
5.7835 | -101.5954
5.7872 | -101.5972
5.7908 | -101.6034
5.7945 | -101.6117
5.7982 | -101.6229
5.8018 | -101.6269
5.8055 | -101.6299
5.8091 | -101.624
5.8128 | -101.6117
5.8165 | -101.5984
5.8201 | -101.5838
5.8238 | -101.5656
5.8274 | -101.5428
5.8311 | -101.533
5.8347 | -101.5363
5.8384 | -101.55
5.8421 | -101.5771
5.8457 | -101.6265
5.8494 | -101.6876
5.853 | -101.7537
5.8567 | -101.8149
5.8604 | -101.8733
5.864 | -101.917
5.8677 | -101.9492
5.8713 | -101.9635
5.875 | -101.952
5.8786 | -101.9253
5.8823 | -101.8805
5.886 | -101.8199
5.8896 | -101.75
5.8933 | -101.6775
5.8969 | -101.6052
5.9006 | -101.543
5.9042 | -101.4972
5.9079 | -101.4609
5.9116 | -101.4465
5.9152 | -101.4531
5.9189 | -101.4806
5.9225 | -101.521
5.9262 | -101.5722
5.9299 | -101.6347
5.9335 | -101.6945
5.9372 | -101.7463
5.9408 | -101.7896
5.9445 | -101.8146
5.9481 | -101.8187
5.9518 | -101.7981
5.9555 | -101.7649
5.9591 | -101.7228
5.9628 | -101.6795
5.9664 | -101.636
5.9701 | -101.6018
5.9738 | -101.5824
5.9774 | -101.5735
5.9811 | -101.5759
5.9847 | -101.595
5.9884 | -101.6235
5.992 | -101.6589
5.9957 | -101.6904
5.9994 | -101.7172
6.003 | -101.7351
6.0067 | -101.7536
6.0103 | -101.7603
6.014 | -101.761
6.0177 | -101.7615
6.0213 | -101.7619
6.025 | -101.7666
6.0286 | -101.7739
6.0323 | -101.786
6.0359 | -101.8105
6.0396 | -101.8468
6.0433 | -101.8882
6.0469 | -101.9322
6.0506 | -101.9769
6.0542 | -102.0212
6.0579 | -102.0709
6.0615 | -102.1006
6.0652 | -102.1202
6.0689 | -102.1311
6.0725 | -102.1275
6.0762 | -102.1096
6.0798 | -102.0831
6.0835 | -102.0476
6.0872 | -102.0106
6.0908 | -101.9788
6.0945 | -101.9507
6.0981 | -101.9288
6.1018 | -101.9206
6.1054 | -101.924
6.1091 | -101.9305
6.1128 | -101.9499
6.1164 | -101.9787
6.1201 | -102.0099
6.1237 | -102.0414
6.1274 | -102.0672
6.1311 | -102.0864
6.1347 | -102.1036
6.1384 | -102.111
6.142 | -102.106
6.1457 | -102.093
6.1493 | -102.0662
6.153 | -102.0346
6.1567 | -101.9934
6.1603 | -101.9446
6.164 | -101.894
6.1676 | -101.8528
6.1713 | -101.8242
6.175 | -101.8007
6.1786 | -101.7917
6.1823 | -101.7946
6.1859 | -101.8056
6.1896 | -101.8255
6.1932 | -101.8424
6.1969 | -101.8592
6.2006 | -101.8808
6.2042 | -101.8965
6.2079 | -101.9082
6.2115 | -101.9225
6.2152 | -101.9359
6.2188 | -101.943
6.2225 | -101.9476
6.2262 | -101.9441
6.2298 | -101.9344
6.2335 | -101.9185
6.2371 | -101.9034
6.2408 | -101.896
6.2445 | -101.8954
6.2481 | -101.9024
6.2518 | -101.9165
6.2554 | -101.9368
6.2591 | -101.9654
6.2627 | -102.0026
6.2664 | -102.0518
6.2701 | -102.1011
6.2737 | -102.1534
6.2774 | -102.2099
6.281 | -102.2662
6.2847 | -102.323
6.2884 | -102.373
6.292 | -102.417
6.2957 | -102.4582
6.2993 | -102.49
6.303 | -102.507
6.3066 | -102.5165
6.3103 | -102.5215
6.314 | -102.5154
6.3176 | -102.5086
6.3213 | -102.4974
6.3249 | -102.4816
6.3286 | -102.4671
6.3323 | -102.4744
6.3359 | -102.4722
6.3396 | -102.4762
6.3432 | -102.488
6.3469 | -102.5035
6.3505 | -102.5122
6.3542 | -102.5143
6.3579 | -102.5163
6.3615 | -102.5164
6.3652 | -102.5127
6.3688 | -102.5039
6.3725 | -102.4924
6.3761 | -102.4721
6.3798 | -102.4508
6.3835 | -102.4263
6.3871 | -102.3915
6.3908 | -102.3575
6.3944 | -102.3312
6.3981 | -102.3124
6.4018 | -102.3043
6.4054 | -102.3031
6.4091 | -102.3165
6.4127 | -102.3352
6.4164 | -102.3654
6.42 | -102.41
6.4237 | -102.4423
6.4274 | -102.4652
6.431 | -102.4835
6.4347 | -102.4909
6.4383 | -102.4844
6.442 | -102.4685
6.4457 | -102.4489
6.4493 | -102.4321
6.453 | -102.3997
6.4566 | -102.3583
6.4603 | -102.3192
6.4639 | -102.2841
6.4676 | -102.2502
6.4713 | -102.2114
6.4749 | -102.1806
6.4786 | -102.1506
6.4822 | -102.1179
6.4859 | -102.0899
6.4896 | -102.0716
6.4932 | -102.0619
6.4969 | -102.0644
6.5005 | -102.0729
6.5042 | -102.098
6.5078 | -102.1317
6.5115 | -102.1756
6.5152 | -102.2238
6.5188 | -102.2648
6.5225 | -102.3109
6.5261 | -102.3535
6.5298 | -102.389
6.5334 | -102.4201
6.5371 | -102.4381
6.5408 | -102.4401
6.5444 | -102.4337
6.5481 | -102.4099
6.5517 | -102.3841
6.5554 | -102.3471
6.5591 | -102.3005
6.5627 | -102.2573
6.5664 | -102.2209
6.57 | -102.2013
6.5737 | -102.1959
6.5773 | -102.2093
6.581 | -102.2442
6.5847 | -102.2959
6.5883 | -102.3558
6.592 | -102.4217
6.5956 | -102.4945
6.5993 | -102.5718
6.603 | -102.6408
6.6066 | -102.6988
6.6103 | -102.7466
6.6139 | -102.7784
6.6176 | -102.8058
6.6212 | -102.822
6.6249 | -102.8253
6.6286 | -102.818
6.6322 | -102.809
6.6359 | -102.8036
6.6395 | -102.7889
6.6432 | -102.7692
6.6469 | -102.7554
6.6505 | -102.7367
6.6542 | -102.7133
6.6578 | -102.6849
6.6615 | -102.6445
6.6651 | -102.6008
6.6688 | -102.555
6.6725 | -102.5046
6.6761 | -102.453
6.6798 | -102.4059
6.6834 | -102.3646
6.6871 | -102.3394
6.6907 | -102.33
6.6944 | -102.3295
6.6981 | -102.3456
6.7017 | -102.3784
6.7054 | -102.4253
6.709 | -102.4819
6.7127 | -102.5511
6.7164 | -102.6186
6.72 | -102.6858
6.7237 | -102.7586
6.7273 | -102.8302
6.731 | -102.8963
6.7346 | -102.9544
6.7383 | -103.0101
6.742 | -103.0516
6.7456 | -103.071
6.7493 | -103.0729
6.7529 | -103.0656
6.7566 | -103.0455
6.7603 | -103.018
6.7639 | -102.9807
6.7676 | -102.934
6.7712 | -102.8882
6.7749 | -102.8489
6.7785 | -102.8066
6.7822 | -102.7679
6.7859 | -102.7326
6.7895 | -102.7097
6.7932 | -102.695
6.7968 | -102.6826
6.8005 | -102.6787
6.8042 | -102.6837
6.8078 | -102.6946
6.8115 | -102.7129
6.8151 | -102.738
6.8188 | -102.7665
6.8224 | -102.788
6.8261 | -102.8075
6.8298 | -102.8239
6.8334 | -102.8332
6.8371 | -102.8305
6.8407 | -102.8208
6.8444 | -102.8101
6.848 | -102.7904
6.8517 | -102.7548
6.8554 | -102.7152
6.859 | -102.6684
6.8627 | -102.6238
6.8663 | -102.5848
6.87 | -102.5543
6.8737 | -102.5283
6.8773 | -102.5052
6.881 | -102.4918
6.8846 | -102.4836
6.8883 | -102.4863
6.8919 | -102.5011
6.8956 | -102.5278
6.8993 | -102.5566
6.9029 | -102.5882
6.9066 | -102.6202
6.9102 | -102.6436
6.9139 | -102.6533
6.9176 | -102.6554
6.9212 | -102.6519
6.9249 | -102.6422
6.9285 | -102.6264
6.9322 | -102.6006
6.9358 | -102.5747
6.9395 | -102.5516
6.9432 | -102.5252
6.9468 | -102.5028
6.9505 | -102.4885
6.9541 | -102.4829
6.9578 | -102.4851
6.9615 | -102.4967
6.9651 | -102.5204
6.9688 | -102.5587
6.9724 | -102.5948
6.9761 | -102.6339
6.9797 | -102.6773
6.9834 | -102.7168
6.9871 | -102.7564
6.9907 | -102.7946
6.9944 | -102.8225
6.998 | -102.8416
7.0017 | -102.8497
7.0053 | -102.8604
7.009 | -102.8628
7.0127 | -102.8599
7.0163 | -102.857
7.02 | -102.8508
7.0236 | -102.8425
7.0273 | -102.8348
7.031 | -102.8279
7.0346 | -102.8203
7.0383 | -102.814
7.0419 | -102.8054
7.0456 | -102.7973
7.0492 | -102.7903
7.0529 | -102.7797
7.0566 | -102.7682
7.0602 | -102.7599
7.0639 | -102.7529
7.0675 | -102.7508
7.0712 | -102.7521
7.0749 | -102.7551
7.0785 | -102.7654
7.0822 | -102.7795
7.0858 | -102.8022
7.0895 | -102.8316
7.0931 | -102.8615
7.0968 | -102.8911
7.1005 | -102.9244
7.1041 | -102.9599
7.1078 | -103.0016
7.1114 | -103.0421
7.1151 | -103.0799
7.1188 | -103.1112
7.1224 | -103.1369
7.1261 | -103.157
7.1297 | -103.1658
7.1334 | -103.1639
7.137 | -103.1556
7.1407 | -103.1375
7.1444 | -103.118
7.148 | -103.0943
7.1517 | -103.0598
7.1553 | -103.0257
7.159 | -102.9929
7.1627 | -102.9629
7.1663 | -102.929
7.17 | -102.9034
7.1736 | -102.8868
7.1773 | -102.8817
7.1809 | -102.8841
7.1846 | -102.8956
7.1883 | -102.9181
7.1919 | -102.9434
7.1956 | -102.9716
7.1992 | -103.003
7.2029 | -103.0363
7.2065 | -103.0669
7.2102 | -103.0985
7.2139 | -103.1271
7.2175 | -103.1422
7.2212 | -103.1524
7.2248 | -103.1578
7.2285 | -103.1527
7.2322 | -103.1386
7.2358 | -103.115
7.2395 | -103.0899
7.2431 | -103.0641
7.2468 | -103.0457
7.2504 | -103.026
7.2541 | -103.0055
7.2578 | -102.9933
7.2614 | -102.9816
7.2651 | -102.9734
7.2687 | -102.967
7.2724 | -102.9689
7.2761 | -102.9654
7.2797 | -102.9639
7.2834 | -102.9691
7.287 | -102.9804
7.2907 | -103.0007
7.2943 | -103.027
7.298 | -103.0611
7.3017 | -103.1067
7.3053 | -103.1521
7.309 | -103.1983
7.3126 | -103.2489
7.3163 | -103.2925
7.32 | -103.3304
7.3236 | -103.3597
7.3273 | -103.3771
7.3309 | -103.3858
7.3346 | -103.3835
7.3382 | -103.3656
7.3419 | -103.331
7.3456 | -103.2837
7.3492 | -103.2264
7.3529 | -103.1582
7.3565 | -103.0849
7.3602 | -103.0108
7.3638 | -102.939
7.3675 | -102.8749
7.3712 | -102.8179
7.3748 | -102.7712
7.3785 | -102.7345
7.3821 | -102.7135
7.3858 | -102.7284
7.3895 | -102.7211
7.3931 | -102.7178
7.3968 | -102.7216
7.4004 | -102.724
7.4041 | -102.719
7.4077 | -102.7166
7.4114 | -102.7117
7.4151 | -102.7022
7.4187 | -102.6857
7.4224 | -102.6587
7.426 | -102.6273
7.4297 | -102.5942
7.4334 | -102.5651
7.437 | -102.536
7.4407 | -102.507
7.4443 | -102.4836
7.448 | -102.4693
7.4516 | -102.4645
7.4553 | -102.4608
7.459 | -102.4684
7.4626 | -102.4862
7.4663 | -102.5163
7.4699 | -102.5589
7.4736 | -102.6067
7.4773 | -102.656
7.4809 | -102.712
7.4846 | -102.7722
7.4882 | -102.833
7.4919 | -102.8925
7.4955 | -102.9508
7.4992 | -103.0113
7.5029 | -103.0717
7.5065 | -103.128
7.5102 | -103.176
7.5138 | -103.2241
7.5175 | -103.2685
7.5211 | -103.3009
7.5248 | -103.3311
7.5285 | -103.3587
7.5321 | -103.3809
7.5358 | -103.4067
7.5394 | -103.4309
7.5431 | -103.4459
7.5468 | -103.4562
7.5504 | -103.4631
7.5541 | -103.4635
7.5577 | -103.4631
7.5614 | -103.4592
7.565 | -103.4519
7.5687 | -103.4453
7.5724 | -103.4372
7.576 | -103.4287
7.5797 | -103.4176
7.5833 | -103.4052
7.587 | -103.4053
7.5907 | -103.406
7.5943 | -103.3969
7.598 | -103.3878
7.6016 | -103.3767
7.6053 | -103.3629
7.6089 | -103.3445
7.6126 | -103.3231
7.6163 | -103.3024
7.6199 | -103.2877
7.6236 | -103.277
7.6272 | -103.2645
7.6309 | -103.2621
7.6346 | -103.2694
7.6382 | -103.2787
7.6419 | -103.2909
7.6455 | -103.3075
7.6492 | -103.3197
7.6528 | -103.3363
7.6565 | -103.3565
7.6602 | -103.3718
7.6638 | -103.3817
7.6675 | -103.3908
7.6711 | -103.3995
7.6748 | -103.4096
7.6784 | -103.4223
7.6821 | -103.4373
7.6858 | -103.4543
7.6894 | -103.4678
7.6931 | -103.4872
7.6967 | -103.5076
7.7004 | -103.5192
7.7041 | -103.5346
7.7077 | -103.5519
7.7114 | -103.5597
7.715 | -103.5595
7.7187 | -103.5541
7.7223 | -103.5464
7.726 | -103.5353
7.7297 | -103.5161
7.7333 | -103.4957
7.737 | -103.48
7.7406 | -103.4665
7.7443 | -103.4494
7.748 | -103.4255
7.7516 | -103.4047
7.7553 | -103.3934
7.7589 | -103.382
7.7626 | -103.3748
7.7662 | -103.3689
7.7699 | -103.364
7.7736 | -103.3628
7.7772 | -103.372
7.7809 | -103.3793
7.7845 | -103.3832
7.7882 | -103.3891
7.7919 | -103.397
7.7955 | -103.401
7.7992 | -103.4056
7.8028 | -103.41
7.8065 | -103.4025
7.8101 | -103.393
7.8138 | -103.3821
7.8175 | -103.3668
7.8211 | -103.3476
7.8248 | -103.3301
7.8284 | -103.3108
7.8321 | -103.2895
7.8357 | -103.2688
7.8394 | -103.2513
7.8431 | -103.2439
7.8467 | -103.2397
7.8504 | -103.2377
7.854 | -103.2402
7.8577 | -103.2466
7.8614 | -103.2629
7.865 | -103.2793
7.8687 | -103.2925
7.8723 | -103.3127
7.876 | -103.3305
7.8796 | -103.3518
7.8833 | -103.3759
7.887 | -103.3986
7.8906 | -103.4162
7.8943 | -103.443
7.8979 | -103.4708
7.9016 | -103.4956
7.9053 | -103.521
7.9089 | -103.5509
7.9126 | -103.5782
7.9162 | -103.5997
7.9199 | -103.629
7.9235 | -103.6571
7.9272 | -103.6758
7.9309 | -103.6905
7.9345 | -103.706
7.9382 | -103.717
7.9418 | -103.7239
7.9455 | -103.7337
7.9492 | -103.7441
7.9528 | -103.7543
7.9565 | -103.7692
7.9601 | -103.7804
7.9638 | -103.7895
7.9674 | -103.8029
7.9711 | -103.8228
7.9748 | -103.8415
7.9784 | -103.8588
7.9821 | -103.8761
7.9857 | -103.8914
7.9894 | -103.9048
7.9931 | -103.9199
7.9967 | -103.9293
8.0004 | -103.9364
8.004 | -103.9449
8.0077 | -103.948
8.0113 | -103.942
8.015 | -103.9352
8.0187 | -103.9238
8.0223 | -103.9042
8.026 | -103.8851
8.0296 | -103.8685
8.0333 | -103.8514
8.0369 | -103.8338
8.0406 | -103.8118
8.0443 | -103.79
8.0479 | -103.7707
8.0516 | -103.751
8.0552 | -103.7254
8.0589 | -103.7049
8.0626 | -103.6886
8.0662 | -103.6734
8.0699 | -103.6613
8.0735 | -103.6562
8.0772 | -103.6535
8.0808 | -103.6581
8.0845 | -103.6733
8.0882 | -103.6934
8.0918 | -103.7167
8.0955 | -103.7462
8.0991 | -103.7809
8.1028 | -103.8219
8.1065 | -103.8666
8.1101 | -103.9112
8.1138 | -103.9542
8.1174 | -103.9969
8.1211 | -104.0405
8.1247 | -104.082
8.1284 | -104.1259
8.1321 | -104.1663
8.1357 | -104.2009
8.1394 | -104.2314
8.143 | -104.2542
8.1467 | -104.2659
8.1503 | -104.2633
8.154 | -104.2616
8.1577 | -104.2613
8.1613 | -104.259
8.165 | -104.2545
8.1686 | -104.2457
8.1723 | -104.237
8.176 | -104.2247
8.1796 | -104.2047
8.1833 | -104.1793
8.1869 | -104.1499
8.1906 | -104.1159
8.1942 | -104.08
8.1979 | -104.0441
8.2016 | -104.0059
8.2052 | -103.9691
8.2089 | -103.9399
8.2125 | -103.9087
8.2162 | -103.8762
8.2199 | -103.8445
8.2235 | -103.8155
8.2272 | -103.7904
8.2308 | -103.7694
8.2345 | -103.7511
8.2381 | -103.7381
8.2418 | -103.7255
8.2455 | -103.7154
8.2491 | -103.7103
8.2528 | -103.7135
8.2564 | -103.7231
8.2601 | -103.7401
8.2638 | -103.7677
8.2674 | -103.7996
8.2711 | -103.8419
8.2747 | -103.8848
8.2784 | -103.929
8.282 | -103.9796
8.2857 | -104.0339
8.2894 | -104.091
8.293 | -104.1417
8.2967 | -104.1963
8.3003 | -104.2417
8.304 | -104.279
8.3077 | -104.3104
8.3113 | -104.3281
8.315 | -104.3361
8.3186 | -104.3324
8.3223 | -104.321
8.3259 | -104.297
8.3296 | -104.2602
8.3333 | -104.2235
8.3369 | -104.1831
8.3406 | -104.1348
8.3442 | -104.0821
8.3479 | -104.0274
8.3515 | -103.9736
8.3552 | -103.9272
8.3589 | -103.8821
8.3625 | -103.8397
8.3662 | -103.7999
8.3698 | -103.7662
8.3735 | -103.7392
8.3772 | -103.7205
8.3808 | -103.7087
8.3845 | -103.6993
8.3881 | -103.6909
8.3918 | -103.6849
8.3954 | -103.6824
8.3991 | -103.6766
8.4028 | -103.6667
8.4064 | -103.6606
8.4101 | -103.6484
8.4137 | -103.6345
8.4174 | -103.618
8.4211 | -103.6031
8.4247 | -103.5858
8.4284 | -103.5655
8.432 | -103.5479
8.4357 | -103.5344
8.4393 | -103.5236
8.443 | -103.5374
8.4467 | -103.5358
8.4503 | -103.5389
8.454 | -103.547
8.4576 | -103.5626
8.4613 | -103.5861
8.4649 | -103.6131
8.4686 | -103.6389
8.4723 | -103.6657
8.4759 | -103.6949
8.4796 | -103.7204
8.4832 | -103.745
8.4869 | -103.7671
8.4906 | -103.7884
8.4942 | -103.8038
8.4979 | -103.817
8.5015 | -103.8256
8.5052 | -103.8292
8.5088 | -103.8339
8.5125 | -103.8415
8.5162 | -103.8465
8.5198 | -103.8513
8.5235 | -103.8636
8.5271 | -103.8786
8.5308 | -103.8915
8.5345 | -103.912
8.5381 | -103.9336
8.5418 | -103.9542
8.5454 | -103.9759
8.5491 | -103.9985
8.5527 | -104.0191
8.5564 | -104.0381
8.5601 | -104.0473
8.5637 | -104.0555
8.5674 | -104.0588
8.571 | -104.0509
8.5747 | -104.0424
8.5784 | -104.0334
8.582 | -104.0194
8.5857 | -104.0019
8.5893 | -103.9836
8.593 | -103.9646
8.5966 | -103.9456
8.6003 | -103.9279
8.604 | -103.9133
8.6076 | -103.9061
8.6113 | -103.9096
8.6149 | -103.9213
8.6186 | -103.9357
8.6223 | -103.9538
8.6259 | -103.98
8.6296 | -104.0185
8.6332 | -104.0635
8.6369 | -104.112
8.6405 | -104.1601
8.6442 | -104.2086
8.6479 | -104.253
8.6515 | -104.2886
8.6552 | -104.3228
8.6588 | -104.3511
8.6625 | -104.3694
8.6661 | -104.3804
8.6698 | -104.3888
8.6735 | -104.3883
8.6771 | -104.3851
8.6808 | -104.3852
8.6844 | -104.3827
8.6881 | -104.3815
8.6918 | -104.3802
8.6954 | -104.382
8.6991 | -104.3927
8.7027 | -104.4028
8.7064 | -104.4068
8.71 | -104.4136
8.7137 | -104.4187
8.7174 | -104.422
8.721 | -104.4266
8.7247 | -104.4324
8.7283 | -104.4254
8.732 | -104.4062
8.7357 | -104.3819
8.7393 | -104.3531
8.743 | -104.3253
8.7466 | -104.2968
8.7503 | -104.265
8.7539 | -104.229
8.7576 | -104.1899
8.7613 | -104.153
8.7649 | -104.1176
8.7686 | -104.0841
8.7722 | -104.0501
8.7759 | -104.0211
8.7796 | -103.9915
8.7832 | -103.9624
8.7869 | -103.9352
8.7905 | -103.9147
8.7942 | -103.9014
8.7978 | -103.89
8.8015 | -103.883
8.8052 | -103.8774
8.8088 | -103.8748
8.8125 | -103.8762
8.8161 | -103.8838
8.8198 | -103.8961
8.8234 | -103.9129
8.8271 | -103.9285
8.8308 | -103.9475
8.8344 | -103.9692
8.8381 | -103.9885
8.8417 | -104.0103
8.8454 | -104.0363
8.8491 | -104.0619
8.8527 | -104.091
8.8564 | -104.1244
8.86 | -104.1563
8.8637 | -104.1873
8.8673 | -104.2155
8.871 | -104.2412
8.8747 | -104.2689
8.8783 | -104.2935
8.882 | -104.3104
8.8856 | -104.3244
8.8893 | -104.3371
8.893 | -104.3464
8.8966 | -104.3546
8.9003 | -104.3629
8.9039 | -104.3705
8.9076 | -104.3793
8.9112 | -104.3828
8.9149 | -104.3844
8.9186 | -104.3832
8.9222 | -104.3798
8.9259 | -104.3741
8.9295 | -104.3722
8.9332 | -104.3662
8.9369 | -104.3551
8.9405 | -104.3513
8.9442 | -104.3449
8.9478 | -104.337
8.9515 | -104.3289
8.9551 | -104.324
8.9588 | -104.321
8.9625 | -104.3195
8.9661 | -104.3241
8.9698 | -104.3409
8.9734 | -104.3494
8.9771 | -104.3561
8.9807 | -104.3685
8.9844 | -104.3841
8.9881 | -104.4058
8.9917 | -104.4281
8.9954 | -104.4514
8.999 | -104.4738
9.0027 | -104.495
9.0064 | -104.5176
9.01 | -104.5413
9.0137 | -104.5656
9.0173 | -104.585
9.021 | -104.5997
9.0246 | -104.6098
9.0283 | -104.617
9.032 | -104.6204
9.0356 | -104.6163
9.0393 | -104.6135
9.0429 | -104.6096
9.0466 | -104.6028
9.0503 | -104.5952
9.0539 | -104.5873
9.0576 | -104.5818
9.0612 | -104.5786
9.0649 | -104.5802
9.0685 | -104.5788
9.0722 | -104.5711
9.0759 | -104.571
9.0795 | -104.5776
9.0832 | -104.5862
9.0868 | -104.5924
9.0905 | -104.5997
9.0942 | -104.6044
9.0978 | -104.606
9.1015 | -104.605
9.1051 | -104.5991
9.1088 | -104.5927
9.1124 | -104.5845
9.1161 | -104.5699
9.1198 | -104.5502
9.1234 | -104.5289
9.1271 | -104.5103
9.1307 | -104.4871
9.1344 | -104.4617
9.138 | -104.438
9.1417 | -104.4109
9.1454 | -104.3865
9.149 | -104.3636
9.1527 | -104.3366
9.1563 | -104.312
9.16 | -104.2916
9.1637 | -104.2728
9.1673 | -104.2546
9.171 | -104.2341
9.1746 | -104.2159
9.1783 | -104.2021
9.1819 | -104.1922
9.1856 | -104.1816
9.1893 | -104.1772
9.1929 | -104.1777
9.1966 | -104.1777
9.2002 | -104.1747
9.2039 | -104.1815
9.2076 | -104.19
9.2112 | -104.1962
9.2149 | -104.199
9.2185 | -104.2019
9.2222 | -104.2063
9.2258 | -104.209
9.2295 | -104.2116
9.2332 | -104.213
9.2368 | -104.2085
9.2405 | -104.2028
9.2441 | -104.1979
9.2478 | -104.19
9.2515 | -104.1837
9.2551 | -104.1786
9.2588 | -104.1746
9.2624 | -104.1708
9.2661 | -104.171
9.2697 | -104.1795
9.2734 | -104.1901
9.2771 | -104.2026
9.2807 | -104.2254
9.2844 | -104.2497
9.288 | -104.2821
9.2917 | -104.3165
9.2953 | -104.3521
9.299 | -104.3909
9.3027 | -104.4325
9.3063 | -104.4794
9.31 | -104.5237
9.3136 | -104.5643
9.3173 | -104.5986
9.321 | -104.6235
9.3246 | -104.6406
9.3283 | -104.6556
9.3319 | -104.6679
9.3356 | -104.6805
9.3392 | -104.6819
9.3429 | -104.6802
9.3466 | -104.6745
9.3502 | -104.6673
9.3539 | -104.6604
9.3575 | -104.6555
9.3612 | -104.6493
9.3649 | -104.6432
9.3685 | -104.6382
9.3722 | -104.6418
9.3758 | -104.6496
9.3795 | -104.651
9.3831 | -104.6614
9.3868 | -104.6719
9.3905 | -104.683
9.3941 | -104.6967
9.3978 | -104.7103
9.4014 | -104.7249
9.4051 | -104.7351
9.4088 | -104.7444
9.4124 | -104.7481
9.4161 | -104.7457
9.4197 | -104.7388
9.4234 | -104.7256
9.427 | -104.7125
9.4307 | -104.6941
9.4344 | -104.6718
9.438 | -104.6511
9.4417 | -104.6296
9.4453 | -104.6057
9.449 | -104.5845
9.4527 | -104.5691
9.4563 | -104.5592
9.46 | -104.549
9.4636 | -104.5372
9.4673 | -104.5291
9.4709 | -104.5201
9.4746 | -104.5159
9.4783 | -104.5154
9.4819 | -104.5132
9.4856 | -104.5131
9.4892 | -104.5111
9.4929 | -104.5058
9.4965 | -104.4946
9.5002 | -104.4872
9.5039 | -104.4791
9.5075 | -104.4702
9.5112 | -104.4581
9.5148 | -104.4493
9.5185 | -104.4446
9.5222 | -104.4355
9.5258 | -104.4291
9.5295 | -104.4243
9.5331 | -104.4241
9.5368 | -104.4286
9.5404 | -104.436
9.5441 | -104.4502
9.5478 | -104.4693
9.5514 | -104.495
9.5551 | -104.5212
9.5587 | -104.5533
9.5624 | -104.5894
9.5661 | -104.6236
9.5697 | -104.6611
9.5734 | -104.7011
9.577 | -104.7407
9.5807 | -104.7793
9.5843 | -104.8192
9.588 | -104.8543
9.5917 | -104.8804
9.5953 | -104.907
9.599 | -104.9271
9.6026 | -104.939
9.6063 | -104.9469
9.6099 | -104.9531
9.6136 | -104.949
9.6173 | -104.9381
9.6209 | -104.9284
9.6246 | -104.9072
9.6282 | -104.8829
9.6319 | -104.8579
9.6356 | -104.8309
9.6392 | -104.8013
9.6429 | -104.7696
9.6465 | -104.7374
9.6502 | -104.7058
9.6538 | -104.6787
9.6575 | -104.652
9.6612 | -104.6298
9.6648 | -104.6087
9.6685 | -104.5904
9.6721 | -104.5812
9.6758 | -104.5705
9.6795 | -104.5595
9.6831 | -104.5544
9.6868 | -104.5507
9.6904 | -104.5461
9.6941 | -104.5483
9.6977 | -104.5492
9.7014 | -104.5481
9.7051 | -104.5501
9.7087 | -104.5528
9.7124 | -104.551
9.716 | -104.5474
9.7197 | -104.5426
9.7234 | -104.5384
9.727 | -104.5357
9.7307 | -104.5311
9.7343 | -104.528
9.738 | -104.524
9.7416 | -104.5205
9.7453 | -104.5156
9.749 | -104.5094
9.7526 | -104.5036
9.7563 | -104.4992
9.7599 | -104.492
9.7636 | -104.4831
9.7673 | -104.482
9.7709 | -104.4821
9.7746 | -104.4812
9.7782 | -104.4792
9.7819 | -104.4786
9.7855 | -104.4829
9.7892 | -104.4897
9.7929 | -104.5002
9.7965 | -104.5103
9.8002 | -104.52
9.8038 | -104.5338
9.8075 | -104.5535
9.8111 | -104.5761
9.8148 | -104.5939
9.8185 | -104.6117
9.8221 | -104.6333
9.8258 | -104.6505
9.8294 | -104.6632
9.8331 | -104.6809
9.8368 | -104.6936
9.8404 | -104.7047
9.8441 | -104.7126
9.8477 | -104.7113
9.8514 | -104.7073
9.855 | -104.7037
9.8587 | -104.6984
9.8624 | -104.6907
9.866 | -104.6826
9.8697 | -104.6729
9.8733 | -104.6678
9.877 | -104.6633
9.8807 | -104.6542
9.8843 | -104.6473
9.888 | -104.6454
9.8916 | -104.6435
9.8953 | -104.642
9.8989 | -104.6427
9.9026 | -104.6455
9.9063 | -104.6508
9.9099 | -104.6578
9.9136 | -104.6664
9.9172 | -104.6796
9.9209 | -104.6957
9.9246 | -104.7117
9.9282 | -104.7279
9.9319 | -104.7423
9.9355 | -104.753
9.9392 | -104.7605
9.9428 | -104.7673
9.9465 | -104.7697
9.9502 | -104.7699
9.9538 | -104.7648
9.9575 | -104.7572
9.9611 | -104.7514
9.9648 | -104.7423
9.9684 | -104.7299
9.9721 | -104.7197
9.9758 | -104.7097
9.9794 | -104.6974
9.9831 | -104.6845
9.9867 | -104.677
9.9904 | -104.6719
9.9941 | -104.6636
9.9977 | -104.6568
10.0014 | -104.6515
10.005 | -104.646
10.0087 | -104.6358
10.0123 | -104.6278
10.016 | -104.6214
10.0197 | -104.6168
10.0233 | -104.6149
10.027 | -104.6155
10.0306 | -104.6133
10.0343 | -104.6119
10.038 | -104.6132
10.0416 | -104.6161
10.0453 | -104.6188
10.0489 | -104.6251
10.0526 | -104.6288
10.0562 | -104.633
10.0599 | -104.6383
10.0636 | -104.6405
10.0672 | -104.648
10.0709 | -104.6602
10.0745 | -104.6734
10.0782 | -104.6898
10.0819 | -104.7087
10.0855 | -104.7309
10.0892 | -104.7525
10.0928 | -104.7763
10.0965 | -104.7972
10.1001 | -104.8195
10.1038 | -104.8443
10.1075 | -104.8672
10.1111 | -104.8868
10.1148 | -104.9056
10.1184 | -104.9241
10.1221 | -104.9399
10.1257 | -104.958
10.1294 | -104.9753
10.1331 | -104.9885
10.1367 | -104.9982
10.1404 | -105.0061
10.144 | -105.0103
10.1477 | -105.0108
10.1514 | -105.0101
10.155 | -105.0076
10.1587 | -105.0025
10.1623 | -104.9934
10.166 | -104.9782
10.1696 | -104.9648
10.1733 | -104.9508
10.177 | -104.9346
10.1806 | -104.9196
10.1843 | -104.9028
10.1879 | -104.8865
10.1916 | -104.8711
10.1953 | -104.8613
10.1989 | -104.8504
10.2026 | -104.8409
10.2062 | -104.8301
10.2099 | -104.8229
10.2135 | -104.8203
10.2172 | -104.8173
10.2209 | -104.8119
10.2245 | -104.8025
10.2282 | -104.7965
10.2318 | -104.7947
10.2355 | -104.7917
10.2392 | -104.7878
10.2428 | -104.7823
10.2465 | -104.7802
10.2501 | -104.7777
10.2538 | -104.7725
10.2574 | -104.7668
10.2611 | -104.7612
10.2648 | -104.7559
10.2684 | -104.7494
10.2721 | -104.7405
10.2757 | -104.7333
10.2794 | -104.7311
10.283 | -104.73
10.2867 | -104.7315
10.2904 | -104.7338
10.294 | -104.7372
10.2977 | -104.7439
10.3013 | -104.7546
10.305 | -104.7709
10.3087 | -104.788
10.3123 | -104.8099
10.316 | -104.837
10.3196 | -104.8591
10.3233 | -104.8824
10.3269 | -104.9097
10.3306 | -104.9364
10.3343 | -104.9636
10.3379 | -104.9893
10.3416 | -105.0161
10.3452 | -105.0369
10.3489 | -105.0571
10.3526 | -105.0762
10.3562 | -105.0948
10.3599 | -105.1062
10.3635 | -105.1153
10.3672 | -105.1197
10.3708 | -105.1278
10.3745 | -105.1365
10.3782 | -105.1395
10.3818 | -105.1389
10.3855 | -105.1434
10.3891 | -105.1503
10.3928 | -105.1531
10.3965 | -105.1565
10.4001 | -105.1628
10.4038 | -105.1755
10.4074 | -105.1881
10.4111 | -105.2011
10.4147 | -105.2133
10.4184 | -105.2257
10.4221 | -105.2383
10.4257 | -105.2483
10.4294 | -105.2567
10.433 | -105.2651
10.4367 | -105.2735
10.4403 | -105.2837
10.444 | -105.2894
10.4477 | -105.2898
10.4513 | -105.2907
10.455 | -105.2892
10.4586 | -105.2837
10.4623 | -105.277
10.466 | -105.2718
10.4696 | -105.2686
10.4733 | -105.2651
10.4769 | -105.2554
10.4806 | -105.2464
10.4842 | -105.2358
10.4879 | -105.2231
10.4916 | -105.2074
10.4952 | -105.1927
10.4989 | -105.1774
10.5025 | -105.1645
10.5062 | -105.1496
10.5099 | -105.134
10.5135 | -105.1191
10.5172 | -105.1068
10.5208 | -105.0894
10.5245 | -105.0709
10.5281 | -105.0573
10.5318 | -105.0437
10.5355 | -105.0259
10.5391 | -105.0083
10.5428 | -104.9992
10.5464 | -104.9913
10.5501 | -104.9845
10.5538 | -104.9914
10.5574 | -104.9859
10.5611 | -104.989
10.5647 | -104.9955
10.5684 | -105.0063
10.572 | -105.0184
10.5757 | -105.0336
10.5794 | -105.0503
10.583 | -105.0686
10.5867 | -105.0871
10.5903 | -105.1065
10.594 | -105.1258
10.5976 | -105.1462
10.6013 | -105.1633
10.605 | -105.1746
10.6086 | -105.1863
10.6123 | -105.1976
10.6159 | -105.2021
10.6196 | -105.2039
10.6233 | -105.202
10.6269 | -105.1957
10.6306 | -105.1902
10.6342 | -105.1805
10.6379 | -105.1658
10.6415 | -105.1463
10.6452 | -105.1275
10.6489 | -105.1068
10.6525 | -105.0866
10.6562 | -105.0641
10.6598 | -105.0426
10.6635 | -105.0224
10.6672 | -104.9991
10.6708 | -104.9755
10.6745 | -104.9541
10.6781 | -104.9388
10.6818 | -104.9205
10.6854 | -104.9066
10.6891 | -104.899
10.6928 | -104.8926
10.6964 | -104.8877
10.7001 | -104.8865
10.7037 | -104.8911
10.7074 | -104.901
10.7111 | -104.9122
10.7147 | -104.9213
10.7184 | -104.9303
10.722 | -104.9436
10.7257 | -104.9584
10.7293 | -104.9733
10.733 | -104.9847
10.7367 | -104.9969
10.7403 | -105.0072
10.744 | -105.0198
10.7476 | -105.0293
10.7513 | -105.0367
10.7549 | -105.0407
10.7586 | -105.0492
10.7623 | -105.0532
10.7659 | -105.0548
10.7696 | -105.0547
10.7732 | -105.0525
10.7769 | -105.0472
10.7806 | -105.0444
10.7842 | -105.0416
10.7879 | -105.0378
10.7915 | -105.0359
10.7952 | -105.0323
10.7988 | -105.0305
10.8025 | -105.0257
10.8062 | -105.0215
10.8098 | -105.0175
10.8135 | -105.0128
10.8171 | -105.0115
10.8208 | -105.0112
10.8245 | -105.0123
10.8281 | -105.0172
10.8318 | -105.0225
10.8354 | -105.0237
10.8391 | -105.0269
10.8427 | -105.0284
10.8464 | -105.026
10.8501 | -105.0259
10.8537 | -105.0223
10.8574 | -105.0199
10.861 | -105.016
10.8647 | -105.0113
10.8684 | -105.0031
10.872 | -105.002
10.8757 | -104.9993
10.8793 | -104.9954
10.883 | -104.9888
10.8866 | -104.9839
10.8903 | -104.9813
10.894 | -104.977
10.8976 | -104.9732
10.9013 | -104.9677
10.9049 | -104.9672
10.9086 | -104.9658
10.9123 | -104.9675
10.9159 | -104.9654
10.9196 | -104.9699
10.9232 | -104.976
10.9269 | -104.984
10.9305 | -104.9966
10.9342 | -105.0129
10.9379 | -105.0279
10.9415 | -105.0449
10.9452 | -105.0624
10.9488 | -105.0816
10.9525 | -105.1078
10.9561 | -105.1358
10.9598 | -105.1629
10.9635 | -105.1884
10.9671 | -105.2141
10.9708 | -105.2428
10.9744 | -105.2694
10.9781 | -105.2934
10.9818 | -105.3167
10.9854 | -105.3357
10.9891 | -105.3521
10.9927 | -105.3691
10.9964 | -105.3846
11 | -105.3929
11.0037 | -105.4011
11.0074 | -105.4044
11.011 | -105.4048
11.0147 | -105.4007
11.0183 | -105.3911
11.022 | -105.3769
11.0257 | -105.3656
11.0293 | -105.3448
11.033 | -105.3158
11.0366 | -105.2861
11.0403 | -105.2567
11.0439 | -105.2272
11.0476 | -105.1984
11.0513 | -105.1715
11.0549 | -105.1463
11.0586 | -105.124
11.0622 | -105.1014
11.0659 | -105.084
11.0696 | -105.0701
11.0732 | -105.0601
11.0769 | -105.0516
11.0805 | -105.0592
11.0842 | -105.0567
11.0878 | -105.0538
11.0915 | -105.0514
11.0952 | -105.0493
11.0988 | -105.0499
11.1025 | -105.051
11.1061 | -105.054
11.1098 | -105.0577
11.1134 | -105.056
11.1171 | -105.0519
11.1208 | -105.0474
11.1244 | -105.0421
11.1281 | -105.0352
11.1317 | -105.0247
11.1354 | -105.0133
11.1391 | -105.0057
11.1427 | -104.9953
11.1464 | -104.9821
11.15 | -104.9738
11.1537 | -104.9661
11.1573 | -104.9572
11.161 | -104.9517
11.1647 | -104.9427
11.1683 | -104.9326
11.172 | -104.9234
11.1756 | -104.9135
11.1793 | -104.9078
11.183 | -104.905
11.1866 | -104.9028
11.1903 | -104.9044
11.1939 | -104.91
11.1976 | -104.9158
11.2012 | -104.9232
11.2049 | -104.936
11.2086 | -104.9501
11.2122 | -104.966
11.2159 | -104.9825
11.2195 | -105.0016
11.2232 | -105.0264
11.2269 | -105.056
11.2305 | -105.0818
11.2342 | -105.1102
11.2378 | -105.1383
11.2415 | -105.1639
11.2451 | -105.1908
11.2488 | -105.2179
11.2525 | -105.2438
11.2561 | -105.2675
11.2598 | -105.2893
11.2634 | -105.3048
11.2671 | -105.3181
11.2707 | -105.328
11.2744 | -105.3316
11.2781 | -105.3358
11.2817 | -105.3366
11.2854 | -105.3348
11.289 | -105.3294
11.2927 | -105.3154
11.2964 | -105.299
11.3 | -105.2809
11.3037 | -105.2582
11.3073 | -105.237
11.311 | -105.2125
11.3146 | -105.1827
11.3183 | -105.1511
11.322 | -105.1237
11.3256 | -105.0999
11.3293 | -105.079
11.3329 | -105.0578
11.3366 | -105.0362
11.3403 | -105.0178
11.3439 | -105.0028
11.3476 | -104.9878
11.3512 | -104.973
11.3549 | -104.9633
11.3585 | -104.9574
11.3622 | -104.9566
11.3659 | -104.9612
11.3695 | -104.9662
11.3732 | -104.9773
11.3768 | -104.9913
11.3805 | -105.011
11.3842 | -105.0318
11.3878 | -105.0548
11.3915 | -105.081
11.3951 | -105.1159
11.3988 | -105.15
11.4024 | -105.1864
11.4061 | -105.2256
11.4098 | -105.2629
11.4134 | -105.2996
11.4171 | -105.3391
11.4207 | -105.3753
11.4244 | -105.4088
11.428 | -105.4447
11.4317 | -105.4789
11.4354 | -105.5121
11.439 | -105.5407
11.4427 | -105.5668
11.4463 | -105.5923
11.45 | -105.6089
11.4537 | -105.6262
11.4573 | -105.6413
11.461 | -105.6498
11.4646 | -105.6522
11.4683 | -105.6537
11.4719 | -105.6534
11.4756 | -105.6509
11.4793 | -105.6433
11.4829 | -105.6326
11.4866 | -105.618
11.4902 | -105.5975
11.4939 | -105.5756
11.4976 | -105.5554
11.5012 | -105.5364
11.5049 | -105.5134
11.5085 | -105.488
11.5122 | -105.4649
11.5158 | -105.4429
11.5195 | -105.419
11.5232 | -105.3943
11.5268 | -105.3669
11.5305 | -105.3354
11.5341 | -105.3065
11.5378 | -105.2815
11.5415 | -105.2566
11.5451 | -105.2295
11.5488 | -105.2027
11.5524 | -105.174
11.5561 | -105.1452
11.5597 | -105.112
11.5634 | -105.0773
11.5671 | -105.0441
11.5707 | -105.0127
11.5744 | -104.9857
11.578 | -104.9618
11.5817 | -104.9381
11.5853 | -104.9176
11.589 | -104.902
11.5927 | -104.8883
11.5963 | -104.8771
11.6 | -104.871
11.6036 | -104.8725
11.6073 | -104.8947
11.611 | -104.9042
11.6146 | -104.9187
11.6183 | -104.9399
11.6219 | -104.9613
11.6256 | -104.9865
11.6292 | -105.0148
11.6329 | -105.043
11.6366 | -105.0712
11.6402 | -105.0941
11.6439 | -105.1196
11.6475 | -105.1468
11.6512 | -105.1742
11.6549 | -105.1977
11.6585 | -105.2213
11.6622 | -105.2414
11.6658 | -105.2556
11.6695 | -105.2653
11.6731 | -105.2757
11.6768 | -105.2788
11.6805 | -105.2791
11.6841 | -105.278
11.6878 | -105.274
11.6914 | -105.2693
11.6951 | -105.2637
11.6988 | -105.258
11.7024 | -105.2452
11.7061 | -105.2281
11.7097 | -105.2113
11.7134 | -105.1944
11.717 | -105.1789
11.7207 | -105.1628
11.7244 | -105.1457
11.728 | -105.1325
11.7317 | -105.1221
11.7353 | -105.1115
11.739 | -105.1012
11.7426 | -105.0896
11.7463 | -105.0805
11.75 | -105.0753
11.7536 | -105.0675
11.7573 | -105.0579
11.7609 | -105.0522
11.7646 | -105.05
11.7683 | -105.0503
11.7719 | -105.0491
11.7756 | -105.0519
11.7792 | -105.0554
11.7829 | -105.0571
11.7865 | -105.0606
11.7902 | -105.0671
11.7939 | -105.0734
11.7975 | -105.0801
11.8012 | -105.0889
11.8048 | -105.0976
11.8085 | -105.1064
11.8122 | -105.1171
11.8158 | -105.1279
11.8195 | -105.1395
11.8231 | -105.1491
11.8268 | -105.1615
11.8304 | -105.1773
11.8341 | -105.1936
11.8378 | -105.2093
11.8414 | -105.2273
11.8451 | -105.2424
11.8487 | -105.2594
11.8524 | -105.2773
11.8561 | -105.2953
11.8597 | -105.3129
11.8634 | -105.3304
11.867 | -105.3482
11.8707 | -105.3676
11.8743 | -105.3891
11.878 | -105.4068
11.8817 | -105.4187
11.8853 | -105.4301
11.889 | -105.4409
11.8926 | -105.4493
11.8963 | -105.4543
11.9 | -105.4547
11.9036 | -105.453
11.9073 | -105.4496
11.9109 | -105.444
11.9146 | -105.4389
11.9182 | -105.4317
11.9219 | -105.4203
11.9256 | -105.4136
11.9292 | -105.407
11.9329 | -105.3952
11.9365 | -105.38
11.9402 | -105.3644
11.9438 | -105.3515
11.9475 | -105.334
11.9512 | -105.32
11.9548 | -105.3102
11.9585 | -105.2955
11.9621 | -105.2793
11.9658 | -105.2654
11.9695 | -105.2544
11.9731 | -105.2465
11.9768 | -105.2343
11.9804 | -105.2211
11.9841 | -105.2107
11.9877 | -105.2007
11.9914 | -105.1908
11.9951 | -105.184
11.9987 | -105.1761
12.0024 | -105.166
12.006 | -105.158
12.0097 | -105.1537
12.0134 | -105.1494
12.017 | -105.1457
12.0207 | -105.1426
12.0243 | -105.1366
12.028 | -105.1314
12.0316 | -105.1289
12.0353 | -105.1266
12.039 | -105.1216
12.0426 | -105.1147
12.0463 | -105.1102
12.0499 | -105.1099
12.0536 | -105.1137
12.0572 | -105.1145
12.0609 | -105.1154
12.0646 | -105.1188
12.0682 | -105.1261
12.0719 | -105.1334
12.0755 | -105.1393
12.0792 | -105.1466
12.0829 | -105.1547
12.0865 | -105.1632
12.0902 | -105.1723
12.0938 | -105.1851
12.0975 | -105.1995
12.1011 | -105.2117
12.1048 | -105.2253
12.1085 | -105.2383
12.1121 | -105.2478
12.1158 | -105.2567
12.1194 | -105.267
12.1231 | -105.2755
12.1268 | -105.2823
12.1304 | -105.2894
12.1341 | -105.2926
12.1377 | -105.2857
12.1414 | -105.2888
12.145 | -105.2886
12.1487 | -105.2819
12.1524 | -105.2752
12.156 | -105.2652
12.1597 | -105.255
12.1633 | -105.2421
12.167 | -105.2264
12.1707 | -105.2095
12.1743 | -105.1912
12.178 | -105.1734
12.1816 | -105.1536
12.1853 | -105.1316
12.1889 | -105.1102
12.1926 | -105.0898
12.1963 | -105.0715
12.1999 | -105.0556
12.2036 | -105.0383
12.2072 | -105.0223
12.2109 | -105.0088
12.2145 | -104.9993
12.2182 | -104.9941
12.2219 | -104.9887
12.2255 | -104.9828
12.2292 | -104.9779
12.2328 | -104.9788
12.2365 | -104.9821
12.2402 | -104.9866
12.2438 | -104.9912
12.2475 | -105.0024
12.2511 | -105.0158
12.2548 | -105.028
12.2584 | -105.0404
12.2621 | -105.0523
12.2658 | -105.0672
12.2694 | -105.0838
12.2731 | -105.0998
12.2767 | -105.1165
12.2804 | -105.1354
12.2841 | -105.1559
12.2877 | -105.1772
12.2914 | -105.1966
12.295 | -105.2144
12.2987 | -105.2297
12.3023 | -105.2427
12.306 | -105.2585
12.3097 | -105.2739
12.3133 | -105.2885
12.317 | -105.3008
12.3206 | -105.3109
12.3243 | -105.3195
12.328 | -105.3308
12.3316 | -105.3432
12.3353 | -105.3554
12.3389 | -105.3638
12.3426 | -105.3691
12.3462 | -105.3763
12.3499 | -105.3834
12.3536 | -105.3887
12.3572 | -105.3931
12.3609 | -105.3975
12.3645 | -105.3987
12.3682 | -105.3964
12.3719 | -105.3957
12.3755 | -105.3935
12.3792 | -105.3886
12.3828 | -105.3813
12.3865 | -105.3741
12.3901 | -105.3715
12.3938 | -105.37
12.3975 | -105.364
12.4011 | -105.3569
12.4048 | -105.3527
12.4084 | -105.3501
12.4121 | -105.3439
12.4157 | -105.3365
12.4194 | -105.3318
12.4231 | -105.3272
12.4267 | -105.3193
12.4304 | -105.3141
12.434 | -105.3094
12.4377 | -105.307
12.4414 | -105.307
12.445 | -105.3038
12.4487 | -105.3005
12.4523 | -105.2993
12.456 | -105.2998
12.4596 | -105.3009
12.4633 | -105.3015
12.467 | -105.3021
12.4706 | -105.3031
12.4743 | -105.3071
12.4779 | -105.3123
12.4816 | -105.3192
12.4853 | -105.3285
12.4889 | -105.3393
12.4926 | -105.347
12.4962 | -105.3585
12.4999 | -105.3718
12.5035 | -105.384
12.5072 | -105.3952
12.5109 | -105.4073
12.5145 | -105.418
12.5182 | -105.4278
12.5218 | -105.4391
12.5255 | -105.4486
12.5292 | -105.4571
12.5328 | -105.4656
12.5365 | -105.4706
12.5401 | -105.4743
12.5438 | -105.4788
12.5474 | -105.4797
12.5511 | -105.4788
12.5548 | -105.4829
12.5584 | -105.4868
12.5621 | -105.4922
12.5657 | -105.4959
12.5694 | -105.4962
12.573 | -105.4947
12.5767 | -105.4971
12.5804 | -105.5003
12.584 | -105.5013
12.5877 | -105.5035
12.5913 | -105.5052
12.595 | -105.5091
12.5987 | -105.5113
12.6023 | -105.5138
12.606 | -105.5183
12.6096 | -105.5205
12.6133 | -105.5187
12.6169 | -105.515
12.6206 | -105.5089
12.6243 | -105.5052
12.6279 | -105.5003
12.6316 | -105.4941
12.6352 | -105.4914
12.6389 | -105.4877
12.6426 | -105.4823
12.6462 | -105.4775
12.6499 | -105.4749
12.6535 | -105.4694
12.6572 | -105.4628
12.6608 | -105.4597
12.6645 | -105.4571
12.6682 | -105.4469
12.6718 | -105.4398
12.6755 | -105.4373
12.6791 | -105.4326
12.6828 | -105.4295
12.6865 | -105.4289
12.6901 | -105.4264
12.6938 | -105.425
12.6974 | -105.4249
12.7011 | -105.4253
12.7047 | -105.4271
12.7084 | -105.4291
12.7121 | -105.4326
12.7157 | -105.4388
12.7194 | -105.4491
12.723 | -105.4611
12.7267 | -105.4703
12.7303 | -105.481
12.734 | -105.4914
12.7377 | -105.4963
12.7413 | -105.5028
12.745 | -105.5079
12.7486 | -105.5135
12.7523 | -105.519
12.756 | -105.5264
12.7596 | -105.5325
12.7633 | -105.5391
12.7669 | -105.5452
12.7706 | -105.5476
12.7742 | -105.5511
12.7779 | -105.5483
12.7816 | -105.5447
12.7852 | -105.5397
12.7889 | -105.5326
12.7925 | -105.525
12.7962 | -105.5196
12.7999 | -105.5082
12.8035 | -105.4958
12.8072 | -105.4797
12.8108 | -105.4637
12.8145 | -105.4462
12.8181 | -105.4246
12.8218 | -105.4053
12.8255 | -105.3898
12.8291 | -105.3762
12.8328 | -105.3621
12.8364 | -105.3484
12.8401 | -105.3321
12.8438 | -105.3137
12.8474 | -105.303
12.8511 | -105.2953
12.8547 | -105.291
12.8584 | -105.2862
12.862 | -105.2782
12.8657 | -105.2717
12.8694 | -105.2677
12.873 | -105.2675
12.8767 | -105.2713
12.8803 | -105.2704
12.884 | -105.2685
12.8876 | -105.2699
12.8913 | -105.2725
12.895 | -105.2736
12.8986 | -105.2746
12.9023 | -105.2757
12.9059 | -105.2747
12.9096 | -105.275
12.9133 | -105.2788
12.9169 | -105.2824
12.9206 | -105.2843
12.9242 | -105.2867
12.9279 | -105.2895
12.9315 | -105.2925
12.9352 | -105.295
12.9389 | -105.2966
12.9425 | -105.2967
12.9462 | -105.2953
12.9498 | -105.2925
12.9535 | -105.2901
12.9572 | -105.291
12.9608 | -105.2923
12.9645 | -105.2961
12.9681 | -105.2999
12.9718 | -105.3043
12.9754 | -105.3073
12.9791 | -105.3082
12.9828 | -105.3076
12.9864 | -105.3148
12.9901 | -105.3214
12.9937 | -105.3297
12.9974 | -105.3357
13.0011 | -105.3412
13.0047 | -105.3483
13.0084 | -105.3538
13.012 | -105.3587
13.0157 | -105.3632
13.0193 | -105.3691
13.023 | -105.3746
13.0267 | -105.3822
13.0303 | -105.3882
13.034 | -105.3905
13.0376 | -105.3911
13.0413 | -105.3896
13.0449 | -105.3888
13.0486 | -105.3889
13.0523 | -105.3866
13.0559 | -105.3811
13.0596 | -105.3756
13.0632 | -105.3682
13.0669 | -105.3646
13.0706 | -105.3629
13.0742 | -105.3581
13.0779 | -105.3535
13.0815 | -105.3538
13.0852 | -105.3531
13.0888 | -105.3544
13.0925 | -105.3591
13.0962 | -105.3639
13.0998 | -105.3723
13.1035 | -105.3813
13.1071 | -105.3927
13.1108 | -105.4059
13.1145 | -105.4214
13.1181 | -105.4417
13.1218 | -105.4601
13.1254 | -105.4762
13.1291 | -105.4956
13.1327 | -105.5184
13.1364 | -105.5405
13.1401 | -105.5602
13.1437 | -105.5768
13.1474 | -105.5926
13.151 | -105.6083
13.1547 | -105.623
13.1584 | -105.6391
13.162 | -105.6547
13.1657 | -105.6672
13.1693 | -105.6805
13.173 | -105.6892
13.1766 | -105.6966
13.1803 | -105.6983
13.184 | -105.6994
13.1876 | -105.6993
13.1913 | -105.6886
13.1949 | -105.6873
13.1986 | -105.6868
13.2022 | -105.6828
13.2059 | -105.6735
13.2096 | -105.6624
13.2132 | -105.6532
13.2169 | -105.6469
13.2205 | -105.6371
13.2242 | -105.6252
13.2279 | -105.6142
13.2315 | -105.6026
13.2352 | -105.5883
13.2388 | -105.5748
13.2425 | -105.5626
13.2461 | -105.5503
13.2498 | -105.533
13.2535 | -105.5182
13.2571 | -105.506
13.2608 | -105.4917
13.2644 | -105.4781
13.2681 | -105.4666
13.2718 | -105.4559
13.2754 | -105.4478
13.2791 | -105.4401
13.2827 | -105.4325
13.2864 | -105.4285
13.29 | -105.4239
13.2937 | -105.4245
13.2974 | -105.4258
13.301 | -105.4255
13.3047 | -105.4256
13.3083 | -105.4245
13.312 | -105.4258
13.3157 | -105.4317
13.3193 | -105.4354
13.323 | -105.438
13.3266 | -105.4433
13.3303 | -105.4507
13.3339 | -105.4575
13.3376 | -105.4611
13.3413 | -105.4669
13.3449 | -105.4709
13.3486 | -105.4711
13.3522 | -105.4731
13.3559 | -105.4769
13.3596 | -105.477
13.3632 | -105.474
13.3669 | -105.4692
13.3705 | -105.4648
13.3742 | -105.4607
13.3778 | -105.4536
13.3815 | -105.4427
13.3852 | -105.4287
13.3888 | -105.4159
13.3925 | -105.3987
13.3961 | -105.3825
13.3998 | -105.3651
13.4034 | -105.3461
13.4071 | -105.3251
13.4108 | -105.3017
13.4144 | -105.2798
13.4181 | -105.2565
13.4217 | -105.2345
13.4254 | -105.2144
13.4291 | -105.1954
13.4327 | -105.1751
13.4364 | -105.1583
13.44 | -105.1464
13.4437 | -105.1396
13.4473 | -105.1321
13.451 | -105.1238
13.4547 | -105.1195
13.4583 | -105.1175
13.462 | -105.1179
13.4656 | -105.1247
13.4693 | -105.1324
13.473 | -105.1426
13.4766 | -105.1567
13.4803 | -105.1719
13.4839 | -105.1912
13.4876 | -105.2114
13.4912 | -105.2306
13.4949 | -105.2539
13.4986 | -105.2767
13.5022 | -105.2992
13.5059 | -105.3214
13.5095 | -105.3411
13.5132 | -105.3598
13.5168 | -105.3752
13.5205 | -105.3869
13.5242 | -105.3951
13.5278 | -105.4058
13.5315 | -105.4163
13.5351 | -105.4223
13.5388 | -105.4293
13.5425 | -105.4347
13.5461 | -105.4373
13.5498 | -105.4377
13.5534 | -105.438
13.5571 | -105.4375
13.5607 | -105.4358
13.5644 | -105.4366
13.5681 | -105.4353
13.5717 | -105.4353
13.5754 | -105.4368
13.579 | -105.4366
13.5827 | -105.4347
13.5864 | -105.4348
13.59 | -105.4362
13.5937 | -105.4403
13.5973 | -105.4438
13.601 | -105.4486
13.6046 | -105.454
13.6083 | -105.4614
13.612 | -105.4691
13.6156 | -105.4744
13.6193 | -105.4788
13.6229 | -105.4827
13.6266 | -105.4886
13.6303 | -105.4926
13.6339 | -105.4937
13.6376 | -105.4933
13.6412 | -105.496
13.6449 | -105.5019
13.6485 | -105.5068
13.6522 | -105.5124
13.6559 | -105.5158
13.6595 | -105.5132
13.6632 | -105.5142
13.6668 | -105.517
13.6705 | -105.5194
13.6742 | -105.5203
13.6778 | -105.5209
13.6815 | -105.5216
13.6851 | -105.5239
13.6888 | -105.5245
13.6924 | -105.5256
13.6961 | -105.5285
13.6998 | -105.5316
13.7034 | -105.5319
13.7071 | -105.533
13.7107 | -105.5347
13.7144 | -105.5356
13.718 | -105.5372
13.7217 | -105.5406
13.7254 | -105.5403
13.729 | -105.5413
13.7327 | -105.5437
13.7363 | -105.5463
13.74 | -105.5497
13.7437 | -105.5548
13.7473 | -105.5603
13.751 | -105.5684
13.7546 | -105.5767
13.7583 | -105.5854
13.7619 | -105.5901
13.7656 | -105.5994
13.7693 | -105.6084
13.7729 | -105.6176
13.7766 | -105.6292
13.7802 | -105.6424
13.7839 | -105.6572
13.7876 | -105.6714
13.7912 | -105.6833
13.7949 | -105.6946
13.7985 | -105.7049
13.8022 | -105.714
13.8058 | -105.7186
13.8095 | -105.7249
13.8132 | -105.7345
13.8168 | -105.7389
13.8205 | -105.741
13.8241 | -105.7424
13.8278 | -105.746
13.8315 | -105.7507
13.8351 | -105.7539
13.8388 | -105.7554
13.8424 | -105.7544
13.8461 | -105.7536
13.8497 | -105.7533
13.8534 | -105.7525
13.8571 | -105.7543
13.8607 | -105.7568
13.8644 | -105.7575
13.868 | -105.7595
13.8717 | -105.7619
13.8753 | -105.7641
13.879 | -105.7658
13.8827 | -105.767
13.8863 | -105.7691
13.89 | -105.7727
13.8936 | -105.7736
13.8973 | -105.7755
13.901 | -105.7813
13.9046 | -105.7848
13.9083 | -105.787
13.9119 | -105.7891
13.9156 | -105.7923
13.9192 | -105.7917
13.9229 | -105.7904
13.9266 | -105.7878
13.9302 | -105.7835
13.9339 | -105.7767
13.9375 | -105.7646
13.9412 | -105.7527
13.9449 | -105.7405
13.9485 | -105.7294
13.9522 | -105.7125
13.9558 | -105.6917
13.9595 | -105.6636
13.9631 | -105.6259
13.9668 | -105.5864
13.9705 | -105.5505
13.9741 | -105.5149
13.9778 | -105.4746
13.9814 | -105.4318
13.9851 | -105.3899
13.9888 | -105.3481
13.9924 | -105.3062
13.9961 | -105.2643
13.9997 | -105.2201
14.0034 | -105.1759
14.007 | -105.1346
14.0107 | -105.0944
14.0144 | -105.0522
14.018 | -105.0141
14.0217 | -104.9766
14.0253 | -104.9375
14.029 | -104.8955
14.0326 | -104.8574
14.0363 | -104.8249
14.04 | -104.795
14.0436 | -104.7727
14.0473 | -104.7534
14.0509 | -104.7379
14.0546 | -104.7285
14.0583 | -104.7241
14.0619 | -104.7255
14.0656 | -104.7311
14.0692 | -104.7389
14.0729 | -104.7524
14.0765 | -104.7699
14.0802 | -104.7902
14.0839 | -104.8149
14.0875 | -104.8412
14.0912 | -104.8676
14.0948 | -104.8969
14.0985 | -104.9239
14.1022 | -104.9521
14.1058 | -104.9818
14.1095 | -105.0129
14.1131 | -105.0417
14.1168 | -105.0684
14.1204 | -105.0975
14.1241 | -105.124
14.1278 | -105.1497
14.1314 | -105.1753
14.1351 | -105.1999
14.1387 | -105.2253
14.1424 | -105.2489
14.1461 | -105.269
14.1497 | -105.2862
14.1534 | -105.304
14.157 | -105.3225
14.1607 | -105.3444
14.1643 | -105.3722
14.168 | -105.3901
14.1717 | -105.4015
14.1753 | -105.4109
14.179 | -105.4205
14.1826 | -105.4285
14.1863 | -105.4368
14.1899 | -105.4446
14.1936 | -105.4554
14.1973 | -105.4658
14.2009 | -105.4777
14.2046 | -105.4882
14.2082 | -105.5003
14.2119 | -105.5112
14.2156 | -105.5203
14.2192 | -105.5296
14.2229 | -105.5412
14.2265 | -105.5556
14.2302 | -105.5719
14.2338 | -105.5858
14.2375 | -105.5986
14.2412 | -105.6105
14.2448 | -105.6191
14.2485 | -105.628
14.2521 | -105.6366
14.2558 | -105.6458
14.2595 | -105.6547
14.2631 | -105.6633
14.2668 | -105.6687
14.2704 | -105.675
14.2741 | -105.6826
14.2777 | -105.6893
14.2814 | -105.6985
14.2851 | -105.7058
14.2887 | -105.7127
14.2924 | -105.7198
14.296 | -105.7276
14.2997 | -105.7344
14.3034 | -105.7422
14.307 | -105.7456
14.3107 | -105.7476
14.3143 | -105.7489
14.318 | -105.7499
14.3216 | -105.7458
14.3253 | -105.7391
14.329 | -105.7378
14.3326 | -105.736
14.3363 | -105.7308
14.3399 | -105.7232
14.3436 | -105.7135
14.3473 | -105.7026
14.3509 | -105.6927
14.3546 | -105.6805
14.3582 | -105.6706
14.3619 | -105.6599
14.3655 | -105.6458
14.3692 | -105.628
14.3729 | -105.613
14.3765 | -105.5981
14.3802 | -105.5836
14.3838 | -105.5682
14.3875 | -105.5538
14.3911 | -105.5385
14.3948 | -105.5236
14.3985 | -105.509
14.4021 | -105.497
14.4058 | -105.4852
14.4094 | -105.4758
14.4131 | -105.4672
14.4168 | -105.4603
14.4204 | -105.4551
14.4241 | -105.4511
14.4277 | -105.4493
14.4314 | -105.4478
14.435 | -105.4489
14.4387 | -105.4504
14.4424 | -105.4533
14.446 | -105.4594
14.4497 | -105.4681
14.4533 | -105.4776
14.457 | -105.4884
14.4607 | -105.5007
14.4643 | -105.5142
14.468 | -105.5285
14.4716 | -105.5446
14.4753 | -105.5656
14.4789 | -105.5886
14.4826 | -105.6123
14.4863 | -105.6361
14.4899 | -105.6553
14.4936 | -105.6777
14.4972 | -105.7018
14.5009 | -105.7243
14.5046 | -105.7477
14.5082 | -105.7698
14.5119 | -105.7924
14.5155 | -105.8135
14.5192 | -105.832
14.5228 | -105.849
14.5265 | -105.8642
14.5302 | -105.8795
14.5338 | -105.8896
14.5375 | -105.8943
14.5411 | -105.8969
14.5448 | -105.8967
14.5484 | -105.8959
14.5521 | -105.8937
14.5558 | -105.8893
14.5594 | -105.882
14.5631 | -105.8757
14.5667 | -105.8666
14.5704 | -105.856
14.5741 | -105.8444
14.5777 | -105.8334
14.5814 | -105.8213
14.585 | -105.8103
14.5887 | -105.7978
14.5923 | -105.7861
14.596 | -105.7776
14.5997 | -105.7692
14.6033 | -105.761
14.607 | -105.7548
14.6106 | -105.7491
14.6143 | -105.7441
14.618 | -105.7395
14.6216 | -105.7357
14.6253 | -105.7324
14.6289 | -105.7344
14.6326 | -105.7372
14.6362 | -105.7372
14.6399 | -105.7383
14.6436 | -105.7417
14.6472 | -105.7466
14.6509 | -105.7512
14.6545 | -105.7553
14.6582 | -105.755
14.6618 | -105.7576
14.6655 | -105.7605
14.6692 | -105.7614
14.6728 | -105.7648
14.6765 | -105.7679
14.6801 | -105.7699
14.6838 | -105.7715
14.6875 | -105.775
14.6911 | -105.779
14.6948 | -105.7838
14.6984 | -105.7908
14.7021 | -105.7948
14.7057 | -105.798
14.7094 | -105.8034
14.7131 | -105.807
14.7167 | -105.8143
14.7204 | -105.8216
14.724 | -105.8317
14.7277 | -105.8424
14.7314 | -105.852
14.735 | -105.8622
14.7387 | -105.8721
14.7423 | -105.8814
14.746 | -105.8903
14.7496 | -105.8996
14.7533 | -105.9093
14.757 | -105.9185
14.7606 | -105.9259
14.7643 | -105.9286
14.7679 | -105.9284
14.7716 | -105.9268
14.7753 | -105.9167
14.7789 | -105.9131
14.7826 | -105.9072
14.7862 | -105.9045
14.7899 | -105.897
14.7935 | -105.8834
14.7972 | -105.8689
14.8009 | -105.8541
14.8045 | -105.8376
14.8082 | -105.8167
14.8118 | -105.793
14.8155 | -105.7692
14.8192 | -105.7401
14.8228 | -105.7097
14.8265 | -105.6817
14.8301 | -105.653
14.8338 | -105.6241
14.8374 | -105.5939
14.8411 | -105.5622
14.8448 | -105.532
14.8484 | -105.5046
14.8521 | -105.4778
14.8557 | -105.4492
14.8594 | -105.4242
14.863 | -105.4017
14.8667 | -105.3819
14.8704 | -105.3676
14.874 | -105.3558
14.8777 | -105.3462
14.8813 | -105.3384
14.885 | -105.3288
14.8887 | -105.3224
14.8923 | -105.3189
14.896 | -105.3169
14.8996 | -105.3134
14.9033 | -105.3133
14.9069 | -105.3191
14.9106 | -105.3242
14.9143 | -105.3304
14.9179 | -105.3379
14.9216 | -105.3446
14.9252 | -105.3513
14.9289 | -105.3592
14.9326 | -105.3677
14.9362 | -105.3755
14.9399 | -105.3826
14.9435 | -105.3891
14.9472 | -105.3965
14.9508 | -105.4066
14.9545 | -105.4172
14.9582 | -105.4261
14.9618 | -105.4339
14.9655 | -105.4416
14.9691 | -105.4475
14.9728 | -105.4555
14.9765 | -105.4661
14.9801 | -105.475
14.9838 | -105.4816
14.9874 | -105.4895
14.9911 | -105.4954
14.9947 | -105.5004
14.9984 | -105.5071
15.0021 | -105.5122
15.0057 | -105.5212
15.0094 | -105.5332
15.013 | -105.5433
15.0167 | -105.5532
15.0203 | -105.5639
15.024 | -105.5741
15.0277 | -105.5866
15.0313 | -105.599
15.035 | -105.6137
15.0386 | -105.6279
15.0423 | -105.6376
15.046 | -105.6467
15.0496 | -105.656
15.0533 | -105.6658
15.0569 | -105.6744
15.0606 | -105.6825
15.0642 | -105.6862
15.0679 | -105.691
15.0716 | -105.7006
15.0752 | -105.7086
15.0789 | -105.7157
15.0825 | -105.7206
15.0862 | -105.7219
15.0899 | -105.7257
15.0935 | -105.7312
15.0972 | -105.7396
15.1008 | -105.75
15.1045 | -105.7576
15.1081 | -105.7633
15.1118 | -105.7714
15.1155 | -105.7851
15.1191 | -105.7939
15.1228 | -105.7997
15.1264 | -105.8073
15.1301 | -105.8176
15.1338 | -105.8249
15.1374 | -105.8344
15.1411 | -105.846
15.1447 | -105.8533
15.1484 | -105.8618
15.152 | -105.8713
15.1557 | -105.8796
15.1594 | -105.8859
15.163 | -105.8902
15.1667 | -105.8929
15.1703 | -105.8917
15.174 | -105.8909
15.1776 | -105.8917
15.1813 | -105.89
15.185 | -105.8895
15.1886 | -105.8853
15.1923 | -105.8797
15.1959 | -105.8761
15.1996 | -105.8719
15.2033 | -105.865
15.2069 | -105.8567
15.2106 | -105.8491
15.2142 | -105.8411
15.2179 | -105.8331
15.2215 | -105.8221
15.2252 | -105.8076
15.2289 | -105.792
15.2325 | -105.7764
15.2362 | -105.7628
15.2398 | -105.7517
15.2435 | -105.7389
15.2472 | -105.725
15.2508 | -105.7127
15.2545 | -105.7002
15.2581 | -105.6883
15.2618 | -105.6775
15.2654 | -105.6656
15.2691 | -105.6561
15.2728 | -105.6475
15.2764 | -105.6368
15.2801 | -105.63
15.2837 | -105.622
15.2874 | -105.6115
15.2911 | -105.6033
15.2947 | -105.597
15.2984 | -105.5931
15.302 | -105.5905
15.3057 | -105.5881
15.3093 | -105.5826
15.313 | -105.575
15.3167 | -105.5708
15.3203 | -105.5629
15.324 | -105.5574
15.3276 | -105.5568
15.3313 | -105.5554
15.3349 | -105.551
15.3386 | -105.5483
15.3423 | -105.5451
15.3459 | -105.5424
15.3496 | -105.5385
15.3532 | -105.5356
15.3569 | -105.5326
15.3606 | -105.53
15.3642 | -105.5293
15.3679 | -105.5261
15.3715 | -105.5226
15.3752 | -105.5201
15.3788 | -105.5199
15.3825 | -105.5204
15.3862 | -105.5172
15.3898 | -105.5195
15.3935 | -105.5248
15.3971 | -105.5278
15.4008 | -105.5307
15.4045 | -105.5285
15.4081 | -105.53
15.4118 | -105.5323
15.4154 | -105.5317
15.4191 | -105.5333
15.4227 | -105.5367
15.4264 | -105.5406
15.4301 | -105.5447
15.4337 | -105.5475
15.4374 | -105.5474
15.441 | -105.5483
15.4447 | -105.5529
15.4484 | -105.559
15.452 | -105.5641
15.4557 | -105.5665
15.4593 | -105.5662
15.463 | -105.5661
15.4666 | -105.5662
15.4703 | -105.5645
15.474 | -105.5618
15.4776 | -105.5645
15.4813 | -105.5673
15.4849 | -105.5689
15.4886 | -105.5698
15.4922 | -105.57
15.4959 | -105.5711
15.4996 | -105.5732
15.5032 | -105.5716
15.5069 | -105.5722
15.5105 | -105.5728
15.5142 | -105.572
15.5179 | -105.5693
15.5215 | -105.5684
15.5252 | -105.5692
15.5288 | -105.5691
15.5325 | -105.5686
15.5361 | -105.5673
15.5398 | -105.5637
15.5435 | -105.565
15.5471 | -105.5629
15.5508 | -105.5594
15.5544 | -105.5558
15.5581 | -105.5539
15.5618 | -105.5533
15.5654 | -105.5541
15.5691 | -105.5554
15.5727 | -105.5553
15.5764 | -105.556
15.58 | -105.5556
15.5837 | -105.5511
15.5874 | -105.5495
15.591 | -105.5527
15.5947 | -105.5546
15.5983 | -105.5556
15.602 | -105.5563
15.6057 | -105.5567
15.6093 | -105.5587
15.613 | -105.5594
15.6166 | -105.5568
15.6203 | -105.5554
15.6239 | -105.5577
15.6276 | -105.5621
15.6313 | -105.5637
15.6349 | -105.5646
15.6386 | -105.567
15.6422 | -105.5692
15.6459 | -105.568
15.6495 | -105.5687
15.6532 | -105.5698
15.6569 | -105.5711
15.6605 | -105.5736
15.6642 | -105.5736
15.6678 | -105.5721
15.6715 | -105.5689
15.6752 | -105.5667
15.6788 | -105.5639
15.6825 | -105.5596
15.6861 | -105.5559
15.6898 | -105.5527
15.6934 | -105.5503
15.6971 | -105.5481
15.7008 | -105.5397
15.7044 | -105.5297
15.7081 | -105.52
15.7117 | -105.511
15.7154 | -105.5036
15.7191 | -105.4954
15.7227 | -105.4864
15.7264 | -105.4811
15.73 | -105.4763
15.7337 | -105.4706
15.7373 | -105.4666
15.741 | -105.4667
15.7447 | -105.4658
15.7483 | -105.4625
15.752 | -105.462
15.7556 | -105.4622
15.7593 | -105.4644
15.763 | -105.4691
15.7666 | -105.4742
15.7703 | -105.4805
15.7739 | -105.4844
15.7776 | -105.491
15.7812 | -105.4948
15.7849 | -105.4993
15.7886 | -105.5054
15.7922 | -105.5132
15.7959 | -105.5212
15.7995 | -105.5273
15.8032 | -105.5313
15.8069 | -105.5404
15.8105 | -105.5483
15.8142 | -105.5538
15.8178 | -105.5595
15.8215 | -105.5666
15.8251 | -105.5686
15.8288 | -105.5682
15.8325 | -105.5591
15.8361 | -105.5561
15.8398 | -105.5512
15.8434 | -105.5445
15.8471 | -105.5389
15.8507 | -105.5316
15.8544 | -105.5214
15.8581 | -105.5082
15.8617 | -105.494
15.8654 | -105.48
15.869 | -105.4634
15.8727 | -105.4487
15.8764 | -105.4321
15.88 | -105.417
15.8837 | -105.4031
15.8873 | -105.3893
15.891 | -105.3789
15.8946 | -105.3665
15.8983 | -105.3553
15.902 | -105.344
15.9056 | -105.3374
15.9093 | -105.3325
15.9129 | -105.3253
15.9166 | -105.3209
15.9203 | -105.319
15.9239 | -105.3198
15.9276 | -105.324
15.9312 | -105.3292
15.9349 | -105.3386
15.9385 | -105.3461
15.9422 | -105.356
15.9459 | -105.3682
15.9495 | -105.3789
15.9532 | -105.3917
15.9568 | -105.402
15.9605 | -105.4127
15.9642 | -105.4271
15.9678 | -105.441
15.9715 | -105.452
15.9751 | -105.4681
15.9788 | -105.4829
15.9824 | -105.5005
15.9861 | -105.5193
15.9898 | -105.5331
15.9934 | -105.5463
15.9971 | -105.5592
16.0007 | -105.5737
16.0044 | -105.5873
16.008 | -105.6022
16.0117 | -105.6129
16.0154 | -105.6185
16.019 | -105.6251
16.0227 | -105.6314
16.0263 | -105.6388
16.03 | -105.6458
16.0337 | -105.6464
16.0373 | -105.6448
16.041 | -105.6442
16.0446 | -105.6429
16.0483 | -105.6381
16.0519 | -105.6355
16.0556 | -105.6321
16.0593 | -105.627
16.0629 | -105.6215
16.0666 | -105.6183
16.0702 | -105.6142
16.0739 | -105.6102
16.0776 | -105.6052
16.0812 | -105.599
16.0849 | -105.5968
16.0885 | -105.5972
16.0922 | -105.599
16.0958 | -105.6008
16.0995 | -105.6027
16.1032 | -105.6027
16.1068 | -105.6054
16.1105 | -105.6086
16.1141 | -105.612
16.1178 | -105.618
16.1214 | -105.6236
16.1251 | -105.6345
16.1288 | -105.6464
16.1324 | -105.6599
16.1361 | -105.67
16.1397 | -105.6819
16.1434 | -105.6957
16.1471 | -105.7096
16.1507 | -105.7226
16.1544 | -105.7345
16.158 | -105.7474
16.1617 | -105.7589
16.1653 | -105.7683
16.169 | -105.7825
16.1727 | -105.7951
16.1763 | -105.8072
16.18 | -105.8182
16.1836 | -105.8283
16.1873 | -105.8364
16.191 | -105.847
16.1946 | -105.8531
16.1983 | -105.8568
16.2019 | -105.8627
16.2056 | -105.8649
16.2092 | -105.864
16.2129 | -105.8589
16.2166 | -105.8522
16.2202 | -105.8433
16.2239 | -105.833
16.2275 | -105.8217
16.2312 | -105.8054
16.2349 | -105.786
16.2385 | -105.7667
16.2422 | -105.7476
16.2458 | -105.7267
16.2495 | -105.7037
16.2531 | -105.6827
16.2568 | -105.6577
16.2605 | -105.6311
16.2641 | -105.6068
16.2678 | -105.5805
16.2714 | -105.551
16.2751 | -105.5252
16.2788 | -105.5
16.2824 | -105.474
16.2861 | -105.4499
16.2897 | -105.427
16.2934 | -105.4061
16.297 | -105.3855
16.3007 | -105.3645
16.3044 | -105.3436
16.308 | -105.3257
16.3117 | -105.3104
16.3153 | -105.2955
16.319 | -105.2822
16.3226 | -105.2715
16.3263 | -105.2633
16.33 | -105.2576
16.3336 | -105.2545
16.3373 | -105.2515
16.3409 | -105.2511
16.3446 | -105.252
16.3483 | -105.2535
16.3519 | -105.2591
16.3556 | -105.266
16.3592 | -105.2853
16.3629 | -105.2924
16.3665 | -105.3039
16.3702 | -105.3183
16.3739 | -105.3328
16.3775 | -105.35
16.3812 | -105.3667
16.3848 | -105.3847
16.3885 | -105.4062
16.3922 | -105.4278
16.3958 | -105.4509
16.3995 | -105.4752
16.4031 | -105.499
16.4068 | -105.5234
16.4104 | -105.5469
16.4141 | -105.5722
16.4178 | -105.5959
16.4214 | -105.621
16.4251 | -105.645
16.4287 | -105.6702
16.4324 | -105.6953
16.4361 | -105.7187
16.4397 | -105.7414
16.4434 | -105.7661
16.447 | -105.7885
16.4507 | -105.8078
16.4543 | -105.8276
16.458 | -105.8467
16.4617 | -105.8643
16.4653 | -105.8825
16.469 | -105.8971
16.4726 | -105.9108
16.4763 | -105.922
16.4799 | -105.9319
16.4836 | -105.9396
16.4873 | -105.9433
16.4909 | -105.9493
16.4946 | -105.9561
16.4982 | -105.9612
16.5019 | -105.964
16.5056 | -105.9664
16.5092 | -105.9668
16.5129 | -105.9676
16.5165 | -105.9683
16.5202 | -105.9681
16.5238 | -105.9661
16.5275 | -105.9614
16.5312 | -105.9568
16.5348 | -105.9564
16.5385 | -105.959
16.5421 | -105.9579
16.5458 | -105.9563
16.5495 | -105.9587
16.5531 | -105.96
16.5568 | -105.9625
16.5604 | -105.9659
16.5641 | -105.9713
16.5677 | -105.9783
16.5714 | -105.9868
16.5751 | -105.9942
16.5787 | -105.9992
16.5824 | -106.0043
16.586 | -106.0094
16.5897 | -106.0175
16.5934 | -106.0259
16.597 | -106.0366
16.6007 | -106.047
16.6043 | -106.0538
16.608 | -106.0624
16.6116 | -106.0734
16.6153 | -106.0829
16.619 | -106.0903
16.6226 | -106.0979
16.6263 | -106.104
16.6299 | -106.1086
16.6336 | -106.1135
16.6372 | -106.118
16.6409 | -106.1203
16.6446 | -106.122
16.6482 | -106.1234
16.6519 | -106.1257
16.6555 | -106.1278
16.6592 | -106.1271
16.6629 | -106.1243
16.6665 | -106.1198
16.6702 | -106.1162
16.6738 | -106.1122
16.6775 | -106.1071
16.6811 | -106.1001
16.6848 | -106.0954
16.6885 | -106.0913
16.6921 | -106.0869
16.6958 | -106.0806
16.6994 | -106.0741
16.7031 | -106.0676
16.7068 | -106.0599
16.7104 | -106.0526
16.7141 | -106.0463
16.7177 | -106.0398
16.7214 | -106.0327
16.725 | -106.0281
16.7287 | -106.0244
16.7324 | -106.0191
16.736 | -106.0144
16.7397 | -106.0098
16.7433 | -106.006
16.747 | -106.0025
16.7507 | -106.0008
16.7543 | -105.9989
16.758 | -105.9941
16.7616 | -105.9905
16.7653 | -105.9919
16.7689 | -105.994
16.7726 | -105.992
16.7763 | -105.9897
16.7799 | -105.988
16.7836 | -105.9873
16.7872 | -105.9827
16.7909 | -105.98
16.7945 | -105.9736
16.7982 | -105.9683
16.8019 | -105.963
16.8055 | -105.9577
16.8092 | -105.9528
16.8128 | -105.9472
16.8165 | -105.9428
16.8202 | -105.9398
16.8238 | -105.9351
16.8275 | -105.9266
16.8311 | -105.9189
16.8348 | -105.9127
16.8384 | -105.9081
16.8421 | -105.9055
16.8458 | -105.903
16.8494 | -105.8993
16.8531 | -105.8959
16.8567 | -105.8938
16.8604 | -105.8931
16.8641 | -105.8945
16.8677 | -105.8987
16.8714 | -105.9024
16.875 | -105.9055
16.8787 | -105.9082
16.8823 | -105.9139
16.886 | -105.9259
16.8897 | -105.9328
16.8933 | -105.9374
16.897 | -105.9421
16.9006 | -105.9508
16.9043 | -105.9598
16.908 | -105.9696
16.9116 | -105.9787
16.9153 | -105.9877
16.9189 | -105.9983
16.9226 | -106.0084
16.9262 | -106.019
16.9299 | -106.0317
16.9336 | -106.0436
16.9372 | -106.0531
16.9409 | -106.0632
16.9445 | -106.0724
16.9482 | -106.0842
16.9519 | -106.0916
16.9555 | -106.101
16.9592 | -106.1104
16.9628 | -106.1206
16.9665 | -106.1301
16.9701 | -106.1398
16.9738 | -106.1517
16.9775 | -106.1623
16.9811 | -106.1727
16.9848 | -106.1813
16.9884 | -106.1884
16.9921 | -106.1933
16.9957 | -106.1969
16.9994 | -106.2024
17.0031 | -106.205
17.0067 | -106.2029
17.0104 | -106.199
17.014 | -106.197
17.0177 | -106.1938
17.0214 | -106.1905
17.025 | -106.186
17.0287 | -106.1819
17.0323 | -106.1778
17.036 | -106.1754
17.0396 | -106.17
17.0433 | -106.1647
17.047 | -106.1617
17.0506 | -106.1585
17.0543 | -106.1569
17.0579 | -106.159
17.0616 | -106.1597
17.0653 | -106.1597
17.0689 | -106.1637
17.0726 | -106.1689
17.0762 | -106.1743
17.0799 | -106.181
17.0835 | -106.1867
17.0872 | -106.1945
17.0909 | -106.2035
17.0945 | -106.2136
17.0982 | -106.225
17.1018 | -106.2343
17.1055 | -106.2441
17.1091 | -106.2563
17.1128 | -106.2677
17.1165 | -106.2805
17.1201 | -106.293
17.1238 | -106.3026
17.1274 | -106.3123
17.1311 | -106.3215
17.1348 | -106.3316
17.1384 | -106.342
17.1421 | -106.3512
17.1457 | -106.3571
17.1494 | -106.3608
17.153 | -106.3624
17.1567 | -106.3626
17.1604 | -106.362
17.164 | -106.3596
17.1677 | -106.3553
17.1713 | -106.3488
17.175 | -106.3398
17.1787 | -106.333
17.1823 | -106.3271
17.186 | -106.3184
17.1896 | -106.3108
17.1933 | -106.301
17.1969 | -106.2912
17.2006 | -106.2775
17.2043 | -106.2634
17.2079 | -106.2475
17.2116 | -106.2321
17.2152 | -106.2141
17.2189 | -106.1966
17.2226 | -106.1815
17.2262 | -106.1667
17.2299 | -106.1506
17.2335 | -106.1339
17.2372 | -106.1185
17.2408 | -106.1019
17.2445 | -106.0848
17.2482 | -106.0666
17.2518 | -106.0504
17.2555 | -106.0342
17.2591 | -106.0169
17.2628 | -105.9972
17.2664 | -105.9766
17.2701 | -105.9535
17.2738 | -105.9334
17.2774 | -105.9126
17.2811 | -105.8895
17.2847 | -105.8682
17.2884 | -105.848
17.2921 | -105.8252
17.2957 | -105.8012
17.2994 | -105.7771
17.303 | -105.7518
17.3067 | -105.728
17.3103 | -105.7027
17.314 | -105.6757
17.3177 | -105.65
17.3213 | -105.6254
17.325 | -105.6004
17.3286 | -105.5772
17.3323 | -105.5551
17.336 | -105.5343
17.3396 | -105.5141
17.3433 | -105.4955
17.3469 | -105.4767
17.3506 | -105.4596
17.3542 | -105.4436
17.3579 | -105.4313
17.3616 | -105.4206
17.3652 | -105.4144
17.3689 | -105.4101
17.3725 | -105.4065
17.3762 | -105.4062
17.3799 | -105.4056
17.3835 | -105.4086
17.3872 | -105.4156
17.3908 | -105.426
17.3945 | -105.4379
17.3981 | -105.4519
17.4018 | -105.4679
17.4055 | -105.4871
17.4091 | -105.5055
17.4128 | -105.5298
17.4164 | -105.5501
17.4201 | -105.5723
17.4238 | -105.593
17.4274 | -105.6109
17.4311 | -105.6302
17.4347 | -105.6504
17.4384 | -105.6683
17.442 | -105.6854
17.4457 | -105.7038
17.4494 | -105.7226
17.453 | -105.7411
17.4567 | -105.7588
17.4603 | -105.774
17.464 | -105.7887
17.4676 | -105.8018
17.4713 | -105.813
17.475 | -105.8241
17.4786 | -105.8359
17.4823 | -105.8471
17.4859 | -105.8549
17.4896 | -105.863
17.4933 | -105.8696
17.4969 | -105.8757
17.5006 | -105.8819
17.5042 | -105.8882
17.5079 | -105.892
17.5115 | -105.8957
17.5152 | -105.8983
17.5189 | -105.9016
17.5225 | -105.9004
17.5262 | -105.8958
17.5298 | -105.8932
17.5335 | -105.8908
17.5372 | -105.886
17.5408 | -105.8813
17.5445 | -105.8753
17.5481 | -105.8697
17.5518 | -105.8649
17.5554 | -105.8585
17.5591 | -105.8532
17.5628 | -105.8496
17.5664 | -105.8425
17.5701 | -105.8349
17.5737 | -105.8292
17.5774 | -105.8219
17.5811 | -105.8126
17.5847 | -105.8065
17.5884 | -105.7995
17.592 | -105.7913
17.5957 | -105.7869
17.5993 | -105.7811
17.603 | -105.774
17.6067 | -105.7681
17.6103 | -105.7627
17.614 | -105.7574
17.6176 | -105.754
17.6213 | -105.7512
17.6249 | -105.7488
17.6286 | -105.7486
17.6323 | -105.7463
17.6359 | -105.7446
17.6396 | -105.7434
17.6432 | -105.7438
17.6469 | -105.7455
17.6506 | -105.7455
17.6542 | -105.7462
17.6579 | -105.7507
17.6615 | -105.7556
17.6652 | -105.7619
17.6688 | -105.7674
17.6725 | -105.7728
17.6762 | -105.7807
17.6798 | -105.7925
17.6835 | -105.805
17.6871 | -105.8161
17.6908 | -105.8281
17.6945 | -105.8402
17.6981 | -105.8513
17.7018 | -105.8636
17.7054 | -105.8734
17.7091 | -105.8852
17.7127 | -105.8971
17.7164 | -105.907
17.7201 | -105.9181
17.7237 | -105.9276
17.7274 | -105.9354
17.731 | -105.9422
17.7347 | -105.9475
17.7384 | -105.9536
17.742 | -105.9583
17.7457 | -105.9607
17.7493 | -105.9633
17.753 | -105.9627
17.7566 | -105.9597
17.7603 | -105.9586
17.764 | -105.9587
17.7676 | -105.9573
17.7713 | -105.9565
17.7749 | -105.9535
17.7786 | -105.9524
17.7822 | -105.9519
17.7859 | -105.9485
17.7896 | -105.9478
17.7932 | -105.9482
17.7969 | -105.9528
17.8005 | -105.9558
17.8042 | -105.9587
17.8079 | -105.9633
17.8115 | -105.9704
17.8152 | -105.9757
17.8188 | -105.9823
17.8225 | -105.9899
17.8261 | -105.9972
17.8298 | -106.0056
17.8335 | -106.0159
17.8371 | -106.0239
17.8408 | -106.036
17.8444 | -106.0492
17.8481 | -106.0605
17.8518 | -106.0741
17.8554 | -106.089
17.8591 | -106.101
17.8627 | -106.1128
17.8664 | -106.1231
17.87 | -106.1339
17.8737 | -106.1454
17.8774 | -106.1577
17.881 | -106.1709
17.8847 | -106.1837
17.8883 | -106.193
17.892 | -106.202
17.8957 | -106.2125
17.8993 | -106.2249
17.903 | -106.238
17.9066 | -106.249
17.9103 | -106.2568
17.9139 | -106.2644
17.9176 | -106.2701
17.9213 | -106.2753
17.9249 | -106.277
17.9286 | -106.28
17.9322 | -106.2816
17.9359 | -106.2797
17.9395 | -106.2799
17.9432 | -106.2763
17.9469 | -106.2747
17.9505 | -106.2727
17.9542 | -106.271
17.9578 | -106.2644
17.9615 | -106.2593
17.9652 | -106.2524
17.9688 | -106.2456
17.9725 | -106.241
17.9761 | -106.2322
17.9798 | -106.2225
17.9834 | -106.215
17.9871 | -106.205
17.9908 | -106.197
17.9944 | -106.1894
17.9981 | -106.1791
18.0017 | -106.1663
18.0054 | -106.1532
18.0091 | -106.1423
18.0127 | -106.1316
18.0164 | -106.1189
18.02 | -106.1062
18.0237 | -106.094
18.0273 | -106.082
18.031 | -106.0716
18.0347 | -106.0619
18.0383 | -106.0514
18.042 | -106.0422
18.0456 | -106.0305
18.0493 | -106.0195
18.053 | -106.0066
18.0566 | -105.9949
18.0603 | -105.9838
18.0639 | -105.9735
18.0676 | -105.9655
18.0712 | -105.961
18.0749 | -105.9574
18.0786 | -105.953
18.0822 | -105.9485
18.0859 | -105.9462
18.0895 | -105.9438
18.0932 | -105.9432
18.0968 | -105.941
18.1005 | -105.9384
18.1042 | -105.9371
18.1078 | -105.9372
18.1115 | -105.9377
18.1151 | -105.9411
18.1188 | -105.946
18.1225 | -105.9513
18.1261 | -105.9564
18.1298 | -105.9611
18.1334 | -105.9658
18.1371 | -105.9702
18.1407 | -105.9739
18.1444 | -105.9797
18.1481 | -105.9859
18.1517 | -105.99
18.1554 | -105.9924
18.159 | -105.9963
18.1627 | -105.9977
18.1664 | -106.0006
18.17 | -106.0046
18.1737 | -106.0087
18.1773 | -106.0126
18.181 | -106.0179
18.1846 | -106.0217
18.1883 | -106.0232
18.192 | -106.0258
18.1956 | -106.0244
18.1993 | -106.0265
18.2029 | -106.0269
18.2066 | -106.0281
18.2103 | -106.0305
18.2139 | -106.0336
18.2176 | -106.0389
18.2212 | -106.0419
18.2249 | -106.0445
18.2285 | -106.0479
18.2322 | -106.0509
18.2359 | -106.0547
18.2395 | -106.0576
18.2432 | -106.0592
18.2468 | -106.0634
18.2505 | -106.0633
18.2541 | -106.063
18.2578 | -106.0653
18.2615 | -106.0704
18.2651 | -106.0741
18.2688 | -106.0754
18.2724 | -106.0762
18.2761 | -106.0764
18.2798 | -106.0752
18.2834 | -106.0751
18.2871 | -106.0739
18.2907 | -106.073
18.2944 | -106.0716
18.298 | -106.0701
18.3017 | -106.069
18.3054 | -106.0686
18.309 | -106.0687
18.3127 | -106.0701
18.3163 | -106.071
18.32 | -106.0718
18.3237 | -106.071
18.3273 | -106.0673
18.331 | -106.065
18.3346 | -106.0639
18.3383 | -106.0643
18.3419 | -106.0657
18.3456 | -106.0662
18.3493 | -106.0674
18.3529 | -106.0691
18.3566 | -106.0722
18.3602 | -106.0782
18.3639 | -106.0843
18.3676 | -106.0881
18.3712 | -106.0926
18.3749 | -106.0969
18.3785 | -106.1026
18.3822 | -106.1097
18.3858 | -106.1176
18.3895 | -106.1278
18.3932 | -106.1373
18.3968 | -106.1463
18.4005 | -106.1566
18.4041 | -106.1662
18.4078 | -106.1756
18.4114 | -106.1853
18.4151 | -106.194
18.4188 | -106.2025
18.4224 | -106.2119
18.4261 | -106.2202
18.4297 | -106.229
18.4334 | -106.2385
18.4371 | -106.2472
18.4407 | -106.2551
18.4444 | -106.2607
18.448 | -106.266
18.4517 | -106.2705
18.4553 | -106.2765
18.459 | -106.2765
18.4627 | -106.2779
18.4663 | -106.2798
18.47 | -106.2749
18.4736 | -106.2723
18.4773 | -106.2664
18.481 | -106.2618
18.4846 | -106.2564
18.4883 | -106.2501
18.4919 | -106.2433
18.4956 | -106.2349
18.4992 | -106.2267
18.5029 | -106.2186
18.5066 | -106.2107
18.5102 | -106.2025
18.5139 | -106.194
18.5175 | -106.1838
18.5212 | -106.1733
18.5249 | -106.1617
18.5285 | -106.1475
18.5322 | -106.1341
18.5358 | -106.1219
18.5395 | -106.1119
18.5431 | -106.1006
18.5468 | -106.0889
18.5505 | -106.079
18.5541 | -106.0698
18.5578 | -106.0592
18.5614 | -106.0493
18.5651 | -106.041
18.5688 | -106.0333
18.5724 | -106.025
18.5761 | -106.0174
18.5797 | -106.0109
18.5834 | -106.0032
18.587 | -105.9988
18.5907 | -105.9973
18.5944 | -105.9938
18.598 | -105.9915
18.6017 | -105.9897
18.6053 | -105.9884
18.609 | -105.9889
18.6126 | -105.9902
18.6163 | -105.9915
18.62 | -105.9947
18.6236 | -105.9987
18.6273 | -106.0046
18.6309 | -106.0122
18.6346 | -106.0189
18.6383 | -106.0238
18.6419 | -106.0305
18.6456 | -106.0399
18.6492 | -106.0476
18.6529 | -106.0538
18.6565 | -106.0611
18.6602 | -106.0717
18.6639 | -106.0841
18.6675 | -106.0954
18.6712 | -106.1063
18.6748 | -106.117
18.6785 | -106.1273
18.6822 | -106.1384
18.6858 | -106.1504
18.6895 | -106.1591
18.6931 | -106.1652
18.6968 | -106.1698
18.7004 | -106.1747
18.7041 | -106.1788
18.7078 | -106.1827
18.7114 | -106.1858
18.7151 | -106.1885
18.7187 | -106.1895
18.7224 | -106.1887
18.7261 | -106.187
18.7297 | -106.1841
18.7334 | -106.1818
18.737 | -106.1788
18.7407 | -106.1763
18.7443 | -106.1735
18.748 | -106.167
18.7517 | -106.1606
18.7553 | -106.1536
18.759 | -106.1461
18.7626 | -106.1386
18.7663 | -106.1276
18.7699 | -106.118
18.7736 | -106.1085
18.7773 | -106.0996
18.7809 | -106.0921
18.7846 | -106.0862
18.7882 | -106.081
18.7919 | -106.0754
18.7956 | -106.0695
18.7992 | -106.0608
18.8029 | -106.0526
18.8065 | -106.0437
18.8102 | -106.0371
18.8138 | -106.0306
18.8175 | -106.0218
18.8212 | -106.013
18.8248 | -106.0063
18.8285 | -106.0013
18.8321 | -105.9958
18.8358 | -105.9908
18.8395 | -105.986
18.8431 | -105.9816
18.8468 | -105.9793
18.8504 | -105.9768
18.8541 | -105.973
18.8577 | -105.9703
18.8614 | -105.9687
18.8651 | -105.9682
18.8687 | -105.969
18.8724 | -105.9704
18.876 | -105.9728
18.8797 | -105.9758
18.8834 | -105.9807
18.887 | -105.9854
18.8907 | -105.9883
18.8943 | -105.9902
18.898 | -105.9954
18.9016 | -106.0041
18.9053 | -106.0162
18.909 | -106.0249
18.9126 | -106.0319
18.9163 | -106.0421
18.9199 | -106.053
18.9236 | -106.0626
18.9272 | -106.0716
18.9309 | -106.079
18.9346 | -106.0868
18.9382 | -106.0958
18.9419 | -106.1053
18.9455 | -106.1148
18.9492 | -106.1226
18.9529 | -106.1299
18.9565 | -106.1391
18.9602 | -106.1481
18.9638 | -106.1566
18.9675 | -106.1629
18.9711 | -106.1692
18.9748 | -106.177
18.9785 | -106.1843
18.9821 | -106.1911
18.9858 | -106.1971
18.9894 | -106.2063
18.9931 | -106.2116
18.9968 | -106.2158
19.0004 | -106.2212
19.0041 | -106.2265
19.0077 | -106.2315
19.0114 | -106.2379
19.015 | -106.2437
19.0187 | -106.2508
19.0224 | -106.257
19.026 | -106.2624
19.0297 | -106.2673
19.0333 | -106.272
19.037 | -106.279
19.0407 | -106.2834
19.0443 | -106.2864
19.048 | -106.2889
19.0516 | -106.2932
19.0553 | -106.2967
19.0589 | -106.2969
19.0626 | -106.2972
19.0663 | -106.3006
19.0699 | -106.3066
19.0736 | -106.3095
19.0772 | -106.3107
19.0809 | -106.3112
19.0845 | -106.3108
19.0882 | -106.3089
19.0919 | -106.3077
19.0955 | -106.3057
19.0992 | -106.3024
19.1028 | -106.2992
19.1065 | -106.2942
19.1102 | -106.2891
19.1138 | -106.2835
19.1175 | -106.2767
19.1211 | -106.2725
19.1248 | -106.2668
19.1284 | -106.2572
19.1321 | -106.2484
19.1358 | -106.237
19.1394 | -106.2244
19.1431 | -106.2118
19.1467 | -106.2005
19.1504 | -106.1883
19.1541 | -106.1761
19.1577 | -106.166
19.1614 | -106.1558
19.165 | -106.1454
19.1687 | -106.1337
19.1723 | -106.1206
19.176 | -106.107
19.1797 | -106.0924
19.1833 | -106.0794
19.187 | -106.0673
19.1906 | -106.0563
19.1943 | -106.0476
19.198 | -106.0382
19.2016 | -106.0295
19.2053 | -106.0197
19.2089 | -106.0102
19.2126 | -106.0011
19.2162 | -105.9918
19.2199 | -105.9826
19.2236 | -105.976
19.2272 | -105.9708
19.2309 | -105.9667
19.2345 | -105.9601
19.2382 | -105.957
19.2418 | -105.9552
19.2455 | -105.9525
19.2492 | -105.95
19.2528 | -105.9478
19.2565 | -105.9479
19.2601 | -105.9471
19.2638 | -105.9476
19.2675 | -105.9493
19.2711 | -105.9525
19.2748 | -105.9551
19.2784 | -105.9572
19.2821 | -105.9575
19.2857 | -105.9566
19.2894 | -105.9586
19.2931 | -105.9628
19.2967 | -105.968
19.3004 | -105.9726
19.304 | -105.9764
19.3077 | -105.9824
19.3114 | -105.9865
19.315 | -105.9889
19.3187 | -105.992
19.3223 | -105.9962
19.326 | -105.9998
19.3296 | -106.0048
19.3333 | -106.0125
19.337 | -106.0196
19.3406 | -106.0249
19.3443 | -106.0291
19.3479 | -106.0353
19.3516 | -106.0409
19.3553 | -106.0437
19.3589 | -106.0466
19.3626 | -106.0504
19.3662 | -106.0543
19.3699 | -106.0575
19.3735 | -106.0595
19.3772 | -106.0608
19.3809 | -106.0607
19.3845 | -106.0607
19.3882 | -106.0601
19.3918 | -106.0566
19.3955 | -106.0531
19.3991 | -106.0481
19.4028 | -106.0428
19.4065 | -106.0356
19.4101 | -106.0328
19.4138 | -106.0295
19.4174 | -106.021
19.4211 | -106.01
19.4248 | -105.9984
19.4284 | -105.9862
19.4321 | -105.9737
19.4357 | -105.9621
19.4394 | -105.9499
19.443 | -105.9357
19.4467 | -105.9218
19.4504 | -105.9047
19.454 | -105.8892
19.4577 | -105.8731
19.4613 | -105.8555
19.465 | -105.8367
19.4687 | -105.8151
19.4723 | -105.7923
19.476 | -105.7718
19.4796 | -105.753
19.4833 | -105.7326
19.4869 | -105.713
19.4906 | -105.6913
19.4943 | -105.6691
19.4979 | -105.6477
19.5016 | -105.6255
19.5052 | -105.603
19.5089 | -105.5816
19.5126 | -105.5583
19.5162 | -105.5343
19.5199 | -105.5071
19.5235 | -105.4815
19.5272 | -105.4592
19.5308 | -105.4356
19.5345 | -105.4119
19.5382 | -105.3895
19.5418 | -105.3663
19.5455 | -105.3432
19.5491 | -105.3212
19.5528 | -105.3017
19.5565 | -105.2831
19.5601 | -105.2661
19.5638 | -105.2513
19.5674 | -105.2367
19.5711 | -105.2198
19.5747 | -105.2053
19.5784 | -105.1928
19.5821 | -105.1818
19.5857 | -105.1694
19.5894 | -105.1597
19.593 | -105.1537
19.5967 | -105.1474
19.6003 | -105.1418
19.604 | -105.1353
19.6077 | -105.1315
19.6113 | -105.1301
19.615 | -105.1293
19.6186 | -105.1292
19.6223 | -105.129
19.626 | -105.1313
19.6296 | -105.1369
19.6333 | -105.143
19.6369 | -105.15
19.6406 | -105.1581
19.6442 | -105.165
19.6479 | -105.1723
19.6516 | -105.1808
19.6552 | -105.19
19.6589 | -105.2005
19.6625 | -105.2118
19.6662 | -105.2214
19.6699 | -105.2314
19.6735 | -105.2409
19.6772 | -105.2531
19.6808 | -105.2665
19.6845 | -105.2791
19.6881 | -105.2885
19.6918 | -105.2982
19.6955 | -105.3104
19.6991 | -105.323
19.7028 | -105.3347
19.7064 | -105.3458
19.7101 | -105.3543
19.7138 | -105.3634
19.7174 | -105.3731
19.7211 | -105.3823
19.7247 | -105.3903
19.7284 | -105.4011
19.732 | -105.4118
19.7357 | -105.4189
19.7394 | -105.428
19.743 | -105.4384
19.7467 | -105.4488
19.7503 | -105.4579
19.754 | -105.4653
19.7576 | -105.4739
19.7613 | -105.4819
19.765 | -105.4858
19.7686 | -105.4923
19.7723 | -105.4992
19.7759 | -105.5027
19.7796 | -105.5059
19.7833 | -105.5101
19.7869 | -105.5182
19.7906 | -105.5257
19.7942 | -105.5313
19.7979 | -105.5384
19.8015 | -105.5449
19.8052 | -105.5505
19.8089 | -105.5532
19.8125 | -105.5587
19.8162 | -105.5606
19.8198 | -105.5596
19.8235 | -105.5612
19.8272 | -105.5634
19.8308 | -105.5647
19.8345 | -105.5644
19.8381 | -105.5645
19.8418 | -105.5649
19.8454 | -105.5645
19.8491 | -105.5616
19.8528 | -105.5601
19.8564 | -105.5577
19.8601 | -105.5545
19.8637 | -105.5519
19.8674 | -105.5503
19.8711 | -105.5487
19.8747 | -105.5457
19.8784 | -105.5412
19.882 | -105.5353
19.8857 | -105.5338
19.8893 | -105.531
19.893 | -105.5275
19.8967 | -105.5249
19.9003 | -105.5223
19.904 | -105.5196
19.9076 | -105.517
19.9113 | -105.5117
19.9149 | -105.5083
19.9186 | -105.5038
19.9223 | -105.5008
19.9259 | -105.5002
19.9296 | -105.4992
19.9332 | -105.4997
19.9369 | -105.5011
19.9406 | -105.5028
19.9442 | -105.5031
19.9479 | -105.5057
19.9515 | -105.5114
19.9552 | -105.5156
19.9588 | -105.5182
19.9625 | -105.52
19.9662 | -105.5239
19.9698 | -105.53
19.9735 | -105.5369
19.9771 | -105.5436
19.9808 | -105.5516
19.9845 | -105.5599
19.9881 | -105.5669
19.9918 | -105.5758
19.9954 | -105.5854
19.9991 | -105.5956
20.0027 | -105.6056
20.0064 | -105.6151
20.0101 | -105.6215
20.0137 | -105.6275
20.0174 | -105.6357
20.021 | -105.6429
20.0247 | -105.6492
20.0284 | -105.6574
20.032 | -105.6681
20.0357 | -105.6787
20.0393 | -105.6875
20.043 | -105.6966
20.0466 | -105.7075
20.0503 | -105.7178
20.054 | -105.7244
20.0576 | -105.7325
20.0613 | -105.7403
20.0649 | -105.7468
20.0686 | -105.753
20.0722 | -105.758
20.0759 | -105.7639
20.0796 | -105.7709
20.0832 | -105.7769
20.0869 | -105.7795
20.0905 | -105.7835
20.0942 | -105.7858
20.0979 | -105.7867
20.1015 | -105.7903
20.1052 | -105.7924
20.1088 | -105.7921
20.1125 | -105.7903
20.1161 | -105.7865
20.1198 | -105.7816
20.1235 | -105.7775
20.1271 | -105.7744
20.1308 | -105.7708
20.1344 | -105.7679
20.1381 | -105.7629
20.1418 | -105.7566
20.1454 | -105.7502
20.1491 | -105.7432
20.1527 | -105.7352
20.1564 | -105.7275
20.16 | -105.7199
20.1637 | -105.7134
20.1674 | -105.7039
20.171 | -105.6929
20.1747 | -105.6832
20.1783 | -105.6743
20.182 | -105.6677
20.1857 | -105.6614
20.1893 | -105.6561
20.193 | -105.6468
20.1966 | -105.6405
20.2003 | -105.6342
20.2039 | -105.6305
20.2076 | -105.6237
20.2113 | -105.6145
20.2149 | -105.6081
20.2186 | -105.5991
20.2222 | -105.5898
20.2259 | -105.5842
20.2295 | -105.5784
20.2332 | -105.5709
20.2369 | -105.5614
20.2405 | -105.5545
20.2442 | -105.5485
20.2478 | -105.5417
20.2515 | -105.5335
20.2552 | -105.5256
20.2588 | -105.5191
20.2625 | -105.5139
20.2661 | -105.5072
20.2698 | -105.5008
20.2734 | -105.4946
20.2771 | -105.4895
20.2808 | -105.4811
20.2844 | -105.473
20.2881 | -105.4674
20.2917 | -105.4632
20.2954 | -105.4603
20.2991 | -105.4572
20.3027 | -105.4545
20.3064 | -105.4495
20.31 | -105.4454
20.3137 | -105.4409
20.3173 | -105.4372
20.321 | -105.4332
20.3247 | -105.4304
20.3283 | -105.4284
20.332 | -105.4269
20.3356 | -105.4241
20.3393 | -105.4238
20.343 | -105.425
20.3466 | -105.4273
20.3503 | -105.428
20.3539 | -105.4295
20.3576 | -105.4319
20.3612 | -105.434
20.3649 | -105.4355
20.3686 | -105.4391
20.3722 | -105.4432
20.3759 | -105.4491
20.3795 | -105.4538
20.3832 | -105.46
20.3868 | -105.464
20.3905 | -105.4681
20.3942 | -105.475
20.3978 | -105.4823
20.4015 | -105.4875
20.4051 | -105.4915
20.4088 | -105.493
20.4125 | -105.4943
20.4161 | -105.4955
20.4198 | -105.4966
20.4234 | -105.4957
20.4271 | -105.496
20.4307 | -105.4977
20.4344 | -105.4976
20.4381 | -105.4956
20.4417 | -105.4938
20.4454 | -105.4913
20.449 | -105.4887
20.4527 | -105.4855
20.4564 | -105.4816
20.46 | -105.4767
20.4637 | -105.47
20.4673 | -105.4607
20.471 | -105.4503
20.4746 | -105.4395
20.4783 | -105.4303
20.482 | -105.4221
20.4856 | -105.4094
20.4893 | -105.396
20.4929 | -105.3819
20.4966 | -105.3692
20.5003 | -105.3576
20.5039 | -105.3436
20.5076 | -105.3316
20.5112 | -105.3206
20.5149 | -105.3059
20.5185 | -105.2937
20.5222 | -105.2839
20.5259 | -105.2746
20.5295 | -105.2644
20.5332 | -105.2538
20.5368 | -105.2463
20.5405 | -105.241
20.5441 | -105.2349
20.5478 | -105.229
20.5515 | -105.2237
20.5551 | -105.2195
20.5588 | -105.2165
20.5624 | -105.2146
20.5661 | -105.2123
20.5698 | -105.214
20.5734 | -105.2157
20.5771 | -105.2167
20.5807 | -105.2251
20.5844 | -105.2273
20.588 | -105.231
20.5917 | -105.2362
20.5954 | -105.2416
20.599 | -105.2489
20.6027 | -105.256
20.6063 | -105.2648
20.61 | -105.2729
20.6137 | -105.2812
20.6173 | -105.2902
20.621 | -105.2987
20.6246 | -105.3095
20.6283 | -105.3222
20.6319 | -105.3323
20.6356 | -105.3402
20.6393 | -105.349
20.6429 | -105.3615
20.6466 | -105.3732
20.6502 | -105.3832
20.6539 | -105.3932
20.6576 | -105.4032
20.6612 | -105.4115
20.6649 | -105.4209
20.6685 | -105.4297
20.6722 | -105.4382
20.6758 | -105.4468
20.6795 | -105.456
20.6832 | -105.4628
20.6868 | -105.469
20.6905 | -105.4749
20.6941 | -105.4811
20.6978 | -105.4878
20.7015 | -105.4958
20.7051 | -105.5004
20.7088 | -105.5036
20.7124 | -105.5072
20.7161 | -105.5119
20.7197 | -105.5152
20.7234 | -105.5156
20.7271 | -105.5153
20.7307 | -105.515
20.7344 | -105.5154
20.738 | -105.516
20.7417 | -105.5145
20.7453 | -105.5123
20.749 | -105.5129
20.7527 | -105.5121
20.7563 | -105.5118
20.76 | -105.5116
20.7636 | -105.5107
20.7673 | -105.5106
20.771 | -105.5075
20.7746 | -105.5041
20.7783 | -105.5042
20.7819 | -105.5064
20.7856 | -105.504
20.7892 | -105.5008
20.7929 | -105.4986
20.7966 | -105.4972
20.8002 | -105.4948
20.8039 | -105.4895
20.8075 | -105.4853
20.8112 | -105.4803
20.8149 | -105.4748
20.8185 | -105.4702
20.8222 | -105.4663
20.8258 | -105.4633
20.8295 | -105.4585
20.8331 | -105.4541
20.8368 | -105.4492
20.8405 | -105.4448
20.8441 | -105.4409
20.8478 | -105.4359
20.8514 | -105.4344
20.8551 | -105.4336
20.8588 | -105.4325
20.8624 | -105.4307
20.8661 | -105.4296
20.8697 | -105.4297
20.8734 | -105.4307
20.877 | -105.4314
20.8807 | -105.4328
20.8844 | -105.4317
20.888 | -105.432
20.8917 | -105.4331
20.8953 | -105.4349
20.899 | -105.4377
20.9026 | -105.4384
20.9063 | -105.4416
20.91 | -105.4465
20.9136 | -105.4501
20.9173 | -105.4554
20.9209 | -105.4593
20.9246 | -105.4633
20.9283 | -105.4673
20.9319 | -105.4738
20.9356 | -105.4799
20.9392 | -105.4839
20.9429 | -105.489
20.9465 | -105.4941
20.9502 | -105.498
20.9539 | -105.505
20.9575 | -105.5122
20.9612 | -105.5172
20.9648 | -105.5208
20.9685 | -105.5255
20.9722 | -105.5295
20.9758 | -105.5311
20.9795 | -105.5332
20.9831 | -105.5345
20.9868 | -105.5334
20.9904 | -105.5299
20.9941 | -105.528
20.9978 | -105.5241
21.0014 | -105.5212
21.0051 | -105.5193
21.0087 | -105.5139
21.0124 | -105.5052
21.016 | -105.4974
21.0197 | -105.4913
21.0234 | -105.4851
21.027 | -105.4759
21.0307 | -105.4653
21.0343 | -105.4512
21.038 | -105.4409
21.0417 | -105.4283
21.0453 | -105.413
21.049 | -105.3982
21.0526 | -105.3824
21.0563 | -105.3681
21.0599 | -105.3542
21.0636 | -105.339
21.0673 | -105.3252
21.0709 | -105.3122
21.0746 | -105.2968
21.0782 | -105.2796
21.0819 | -105.267
21.0856 | -105.2556
21.0892 | -105.2418
21.0929 | -105.2282
21.0965 | -105.2176
21.1002 | -105.209
21.1038 | -105.2008
21.1075 | -105.2017
21.1112 | -105.1964
21.1148 | -105.1909
21.1185 | -105.1863
21.1221 | -105.1802
21.1258 | -105.1756
21.1295 | -105.1727
21.1331 | -105.1698
21.1368 | -105.1698
21.1404 | -105.1714
21.1441 | -105.1737
21.1477 | -105.1778
21.1514 | -105.177
21.1551 | -105.1757
21.1587 | -105.1779
21.1624 | -105.1818
21.166 | -105.1856
21.1697 | -105.1891
21.1733 | -105.1933
21.177 | -105.1955
21.1807 | -105.2004
21.1843 | -105.2084
21.188 | -105.2115
21.1916 | -105.2153
21.1953 | -105.2184
21.199 | -105.2205
21.2026 | -105.225
21.2063 | -105.2275
21.2099 | -105.2299
21.2136 | -105.2322
21.2172 | -105.2322
21.2209 | -105.2319
21.2246 | -105.2283
21.2282 | -105.2245
21.2319 | -105.2213
21.2355 | -105.2163
21.2392 | -105.2113
21.2429 | -105.2066
21.2465 | -105.2015
21.2502 | -105.1962
21.2538 | -105.1881
21.2575 | -105.18
21.2611 | -105.1699
21.2648 | -105.1594
21.2685 | -105.1496
21.2721 | -105.1375
21.2758 | -105.1271
21.2794 | -105.1169
21.2831 | -105.1051
21.2868 | -105.0947
21.2904 | -105.0847
21.2941 | -105.0744
21.2977 | -105.0665
21.3014 | -105.0572
21.305 | -105.045
21.3087 | -105.0355
21.3124 | -105.0279
21.316 | -105.0189
21.3197 | -105.0094
21.3233 | -105.0011
21.327 | -104.9948
21.3307 | -104.9888
21.3343 | -104.9844
21.338 | -104.9768
21.3416 | -104.9723
21.3453 | -104.9677
21.3489 | -104.9615
21.3526 | -104.955
21.3563 | -104.9489
21.3599 | -104.9431
21.3636 | -104.9375
21.3672 | -104.9327
21.3709 | -104.9278
21.3745 | -104.9248
21.3782 | -104.9218
21.3819 | -104.917
21.3855 | -104.9107
21.3892 | -104.9062
21.3928 | -104.9035
21.3965 | -104.8989
21.4002 | -104.8949
21.4038 | -104.8906
21.4075 | -104.8859
21.4111 | -104.8813
21.4148 | -104.8777
21.4184 | -104.875
21.4221 | -104.8713
21.4258 | -104.8695
21.4294 | -104.8678
21.4331 | -104.8651
21.4367 | -104.8622
21.4404 | -104.861
21.4441 | -104.8593
21.4477 | -104.8578
21.4514 | -104.8565
21.455 | -104.8578
21.4587 | -104.8597
21.4623 | -104.86
21.466 | -104.86
21.4697 | -104.8612
21.4733 | -104.8629
21.477 | -104.864
21.4806 | -104.867
21.4843 | -104.8675
21.488 | -104.8666
21.4916 | -104.8675
21.4953 | -104.8643
21.4989 | -104.8628
21.5026 | -104.8619
21.5062 | -104.8575
21.5099 | -104.8533
21.5136 | -104.8494
21.5172 | -104.8446
21.5209 | -104.8409
21.5245 | -104.8353
21.5282 | -104.8311
21.5318 | -104.8276
21.5355 | -104.8179
21.5392 | -104.8086
21.5428 | -104.8005
21.5465 | -104.7886
21.5501 | -104.776
21.5538 | -104.7645
21.5575 | -104.7513
21.5611 | -104.7357
21.5648 | -104.7221
21.5684 | -104.7071
21.5721 | -104.6908
21.5757 | -104.6756
21.5794 | -104.6579
21.5831 | -104.6376
21.5867 | -104.6172
21.5904 | -104.5962
21.594 | -104.5762
21.5977 | -104.5536
21.6014 | -104.5275
21.605 | -104.5011
21.6087 | -104.475
21.6123 | -104.4495
21.616 | -104.4241
21.6196 | -104.3995
21.6233 | -104.3751
21.627 | -104.3479
21.6306 | -104.3231
21.6343 | -104.2981
21.6379 | -104.2737
21.6416 | -104.2496
21.6453 | -104.2281
21.6489 | -104.2063
21.6526 | -104.1859
21.6562 | -104.1678
21.6599 | -104.1519
21.6635 | -104.1348
21.6672 | -104.1186
21.6709 | -104.1043
21.6745 | -104.0918
21.6782 | -104.0803
21.6818 | -104.0689
21.6855 | -104.0571
21.6891 | -104.0485
21.6928 | -104.0437
21.6965 | -104.0382
21.7001 | -104.0318
21.7038 | -104.0266
21.7074 | -104.0231
21.7111 | -104.0215
21.7148 | -104.0198
21.7184 | -104.0197
21.7221 | -104.0212
21.7257 | -104.0232
21.7294 | -104.0258
21.733 | -104.0291
21.7367 | -104.0334
21.7404 | -104.0382
21.744 | -104.0447
21.7477 | -104.0498
21.7513 | -104.0571
21.755 | -104.0667
21.7587 | -104.0775
21.7623 | -104.0887
21.766 | -104.0996
21.7696 | -104.1099
21.7733 | -104.122
21.7769 | -104.1355
21.7806 | -104.1488
21.7843 | -104.1634
21.7879 | -104.1793
21.7916 | -104.196
21.7952 | -104.2116
21.7989 | -104.2293
21.8026 | -104.2495
21.8062 | -104.2671
21.8099 | -104.2851
21.8135 | -104.3047
21.8172 | -104.3235
21.8208 | -104.3418
21.8245 | -104.3602
21.8282 | -104.3777
21.8318 | -104.3973
21.8355 | -104.4174
21.8391 | -104.4357
21.8428 | -104.4557
21.8464 | -104.4757
21.8501 | -104.4929
21.8538 | -104.5111
21.8574 | -104.5276
21.8611 | -104.5436
21.8647 | -104.5598
21.8684 | -104.576
21.8721 | -104.5914
21.8757 | -104.6059
21.8794 | -104.6177
21.883 | -104.6305
21.8867 | -104.6435
21.8903 | -104.6549
21.894 | -104.6647
21.8977 | -104.6732
21.9013 | -104.6806
21.905 | -104.6868
21.9086 | -104.6918
21.9123 | -104.6963
21.916 | -104.6989
21.9196 | -104.6976
21.9233 | -104.6955
21.9269 | -104.6904
21.9306 | -104.6802
21.9342 | -104.6704
21.9379 | -104.6568
21.9416 | -104.6409
21.9452 | -104.6251
21.9489 | -104.6061
21.9525 | -104.5802
21.9562 | -104.5546
21.9599 | -104.5276
21.9635 | -104.5009
21.9672 | -104.47
21.9708 | -104.4339
21.9745 | -104.3992
21.9781 | -104.3653
21.9818 | -104.3294
21.9855 | -104.2891
21.9891 | -104.2505
21.9928 | -104.2108
21.9964 | -104.1718
22.0001 | -104.1309
22.0038 | -104.0871
22.0074 | -104.0452
22.0111 | -104.0022
22.0147 | -103.96
22.0184 | -103.9181
22.022 | -103.8774
22.0257 | -103.837
22.0294 | -103.7954
22.033 | -103.7534
22.0367 | -103.7141
22.0403 | -103.6773
22.044 | -103.6426
22.0476 | -103.6089
22.0513 | -103.5771
22.055 | -103.5472
22.0586 | -103.5194
22.0623 | -103.4931
22.0659 | -103.4683
22.0696 | -103.4461
22.0733 | -103.4289
22.0769 | -103.4141
22.0806 | -103.4026
22.0842 | -103.3948
22.0879 | -103.3894
22.0915 | -103.3876
22.0952 | -103.3873
22.0989 | -103.3902
22.1025 | -103.3956
22.1062 | -103.4046
22.1098 | -103.4167
22.1135 | -103.4322
22.1172 | -103.4517
22.1208 | -103.4723
22.1245 | -103.4957
22.1281 | -103.5214
22.1318 | -103.5488
22.1354 | -103.5787
22.1391 | -103.6103
22.1428 | -103.6439
22.1464 | -103.6769
22.1501 | -103.7121
22.1537 | -103.7489
22.1574 | -103.7843
22.161 | -103.8214
22.1647 | -103.8596
22.1684 | -103.8974
22.172 | -103.9353
22.1757 | -103.9725
22.1793 | -104.0092
22.183 | -104.0445
22.1867 | -104.0789
22.1903 | -104.1126
22.194 | -104.1453
22.1976 | -104.1745
22.2013 | -104.2018
22.2049 | -104.23
22.2086 | -104.2579
22.2123 | -104.2844
22.2159 | -104.3089
22.2196 | -104.3328
22.2232 | -104.3563
22.2269 | -104.3774
22.2306 | -104.3988
22.2342 | -104.4197
22.2379 | -104.4385
22.2415 | -104.4567
22.2452 | -104.4755
22.2488 | -104.4942
22.2525 | -104.5092
22.2562 | -104.5243
22.2598 | -104.5396
22.2635 | -104.5533
22.2671 | -104.5691
22.2708 | -104.5817
22.2745 | -104.5948
22.2781 | -104.6061
22.2818 | -104.6149
22.2854 | -104.6267
22.2891 | -104.6367
22.2927 | -104.6455
22.2964 | -104.6551
22.3001 | -104.6666
22.3037 | -104.6756
22.3074 | -104.686
22.311 | -104.6954
22.3147 | -104.7075
22.3183 | -104.7186
22.322 | -104.73
22.3257 | -104.7446
22.3293 | -104.757
22.333 | -104.7688
22.3366 | -104.7773
22.3403 | -104.788
22.344 | -104.7999
22.3476 | -104.8133
22.3513 | -104.8231
22.3549 | -104.8317
22.3586 | -104.8402
22.3622 | -104.8489
22.3659 | -104.8583
22.3696 | -104.866
22.3732 | -104.8734
22.3769 | -104.8787
22.3805 | -104.8841
22.3842 | -104.8888
22.3879 | -104.8899
22.3915 | -104.8919
22.3952 | -104.8917
22.3988 | -104.8899
22.4025 | -104.8864
22.4061 | -104.8834
22.4098 | -104.8799
22.4135 | -104.8739
22.4171 | -104.869
22.4208 | -104.8645
22.4244 | -104.8574
22.4281 | -104.8474
22.4318 | -104.8363
22.4354 | -104.8258
22.4391 | -104.8148
22.4427 | -104.8036
22.4464 | -104.792
22.45 | -104.7785
22.4537 | -104.7656
22.4574 | -104.7531
22.461 | -104.743
22.4647 | -104.7327
22.4683 | -104.7234
22.472 | -104.7158
22.4757 | -104.7058
22.4793 | -104.6956
22.483 | -104.6851
22.4866 | -104.6778
22.4903 | -104.6711
22.4939 | -104.6627
22.4976 | -104.6564
22.5013 | -104.6538
22.5049 | -104.6522
22.5086 | -104.6496
22.5122 | -104.6484
22.5159 | -104.6498
22.5195 | -104.6504
22.5232 | -104.6517
22.5269 | -104.6517
22.5305 | -104.6525
22.5342 | -104.653
22.5378 | -104.6554
22.5415 | -104.6601
22.5452 | -104.6651
22.5488 | -104.6693
22.5525 | -104.6728
22.5561 | -104.6759
22.5598 | -104.6803
22.5634 | -104.6864
22.5671 | -104.691
22.5708 | -104.6964
22.5744 | -104.7007
22.5781 | -104.7046
22.5817 | -104.7095
22.5854 | -104.7133
22.5891 | -104.7152
22.5927 | -104.7151
22.5964 | -104.7148
22.6 | -104.718
22.6037 | -104.7194
22.6073 | -104.72
22.611 | -104.7185
22.6147 | -104.7166
22.6183 | -104.7152
22.622 | -104.7125
22.6256 | -104.7088
22.6293 | -104.7047
22.633 | -104.6974
22.6366 | -104.69
22.6403 | -104.6816
22.6439 | -104.6708
22.6476 | -104.6609
22.6512 | -104.6515
22.6549 | -104.6425
22.6586 | -104.6304
22.6622 | -104.6177
22.6659 | -104.6049
22.6695 | -104.5939
22.6732 | -104.5833
22.6768 | -104.5706
22.6805 | -104.557
22.6842 | -104.5422
22.6878 | -104.5274
22.6915 | -104.5139
22.6951 | -104.5012
22.6988 | -104.4894
22.7025 | -104.4783
22.7061 | -104.467
22.7098 | -104.4558
22.7134 | -104.446
22.7171 | -104.4369
22.7207 | -104.4261
22.7244 | -104.4175
22.7281 | -104.411
22.7317 | -104.407
22.7354 | -104.404
22.739 | -104.4
22.7427 | -104.3964
22.7464 | -104.3941
22.75 | -104.3948
22.7537 | -104.3967
22.7573 | -104.4001
22.761 | -104.4028
22.7646 | -104.4056
22.7683 | -104.4119
22.772 | -104.421
22.7756 | -104.4304
22.7793 | -104.4391
22.7829 | -104.4467
22.7866 | -104.4552
22.7903 | -104.4647
22.7939 | -104.4754
22.7976 | -104.4884
22.8012 | -104.5011
22.8049 | -104.5149
22.8085 | -104.527
22.8122 | -104.5415
22.8159 | -104.5578
22.8195 | -104.5733
22.8232 | -104.5876
22.8268 | -104.6022
22.8305 | -104.6171
22.8341 | -104.6328
22.8378 | -104.6474
22.8415 | -104.6629
22.8451 | -104.676
22.8488 | -104.6886
22.8524 | -104.7014
22.8561 | -104.7152
22.8598 | -104.7275
22.8634 | -104.7379
22.8671 | -104.7483
22.8707 | -104.7579
22.8744 | -104.7676
22.878 | -104.7777
22.8817 | -104.7857
22.8854 | -104.7917
22.889 | -104.7975
22.8927 | -104.8039
22.8963 | -104.8098
22.9 | -104.8144
22.9037 | -104.8175
22.9073 | -104.8195
22.911 | -104.8228
22.9146 | -104.8264
22.9183 | -104.8288
22.9219 | -104.8307
22.9256 | -104.8319
22.9293 | -104.8319
22.9329 | -104.8317
22.9366 | -104.8302
22.9402 | -104.8283
22.9439 | -104.8282
22.9476 | -104.8273
22.9512 | -104.8268
22.9549 | -104.8253
22.9585 | -104.8243
22.9622 | -104.8222
22.9658 | -104.8204
22.9695 | -104.8185
22.9732 | -104.8152
22.9768 | -104.8122
22.9805 | -104.8108
22.9841 | -104.8077
22.9878 | -104.8016
22.9914 | -104.7965
22.9951 | -104.7905
22.9988 | -104.7858
23.0024 | -104.7821
23.0061 | -104.7783
23.0097 | -104.7735
23.0134 | -104.7686
23.0171 | -104.7632
23.0207 | -104.7573
23.0244 | -104.7535
23.028 | -104.7469
23.0317 | -104.7405
23.0353 | -104.7358
23.039 | -104.7297
23.0427 | -104.7235
23.0463 | -104.718
23.05 | -104.7133
23.0536 | -104.7058
23.0573 | -104.6986
23.061 | -104.6928
23.0646 | -104.6875
23.0683 | -104.6816
23.0719 | -104.6746
23.0756 | -104.6683
23.0792 | -104.6624
23.0829 | -104.6554
23.0866 | -104.6495
23.0902 | -104.6452
23.0939 | -104.6394
23.0975 | -104.6331
23.1012 | -104.6274
23.1049 | -104.6228
23.1085 | -104.6184
23.1122 | -104.6157
23.1158 | -104.6127
23.1195 | -104.6086
23.1231 | -104.6046
23.1268 | -104.602
23.1305 | -104.5988
23.1341 | -104.5959
23.1378 | -104.5896
23.1414 | -104.5837
23.1451 | -104.579
23.1487 | -104.5735
23.1524 | -104.5679
23.1561 | -104.5655
23.1597 | -104.5609
23.1634 | -104.5564
23.167 | -104.5516
23.1707 | -104.55
23.1744 | -104.5496
23.178 | -104.547
23.1817 | -104.5451
23.1853 | -104.5432
23.189 | -104.5407
23.1926 | -104.5367
23.1963 | -104.5344
23.2 | -104.5328
23.2036 | -104.5296
23.2073 | -104.5272
23.2109 | -104.5267
23.2146 | -104.527
23.2183 | -104.5281
23.2219 | -104.5282
23.2256 | -104.5265
23.2292 | -104.5234
23.2329 | -104.5216
23.2365 | -104.5219
23.2402 | -104.5212
23.2439 | -104.52
23.2475 | -104.5191
23.2512 | -104.5176
23.2548 | -104.516
23.2585 | -104.5152
23.2622 | -104.5148
23.2658 | -104.5142
23.2695 | -104.5121
23.2731 | -104.5098
23.2768 | -104.5082
23.2804 | -104.5053
23.2841 | -104.5026
23.2878 | -104.5019
23.2914 | -104.5017
23.2951 | -104.4985
23.2987 | -104.4952
23.3024 | -104.4932
23.306 | -104.4905
23.3097 | -104.4874
23.3134 | -104.4851
23.317 | -104.4832
23.3207 | -104.48
23.3243 | -104.4767
23.328 | -104.4718
23.3317 | -104.4654
23.3353 | -104.4616
23.339 | -104.4579
23.3426 | -104.4532
23.3463 | -104.449
23.3499 | -104.4448
23.3536 | -104.438
23.3573 | -104.4335
23.3609 | -104.4307
23.3646 | -104.4272
23.3682 | -104.4241
23.3719 | -104.4196
23.3756 | -104.4166
23.3792 | -104.4125
23.3829 | -104.4063
23.3865 | -104.4016
23.3902 | -104.3976
23.3938 | -104.3939
23.3975 | -104.3917
23.4012 | -104.3901
23.4048 | -104.3877
23.4085 | -104.3861
23.4121 | -104.3845
23.4158 | -104.3847
23.4195 | -104.3843
23.4231 | -104.3832
23.4268 | -104.383
23.4304 | -104.3833
23.4341 | -104.3852
23.4377 | -104.3887
23.4414 | -104.3903
23.4451 | -104.3918
23.4487 | -104.3945
23.4524 | -104.3967
23.456 | -104.3981
23.4597 | -104.3995
23.4633 | -104.4015
23.467 | -104.4063
23.4707 | -104.4132
23.4743 | -104.4185
23.478 | -104.4235
23.4816 | -104.4287
23.4853 | -104.4321
23.489 | -104.437
23.4926 | -104.4432
23.4963 | -104.4496
23.4999 | -104.4538
23.5036 | -104.4587
23.5072 | -104.4642
23.5109 | -104.4674
23.5146 | -104.4709
23.5182 | -104.4757
23.5219 | -104.4812
23.5255 | -104.4848
23.5292 | -104.4884
23.5329 | -104.4922
23.5365 | -104.4956
23.5402 | -104.4973
23.5438 | -104.4964
23.5475 | -104.4965
23.5511 | -104.4974
23.5548 | -104.4982
23.5585 | -104.4994
23.5621 | -104.5004
23.5658 | -104.5
23.5694 | -104.4986
23.5731 | -104.4975
23.5768 | -104.4962
23.5804 | -104.4941
23.5841 | -104.4928
23.5877 | -104.4916
23.5914 | -104.4907
23.595 | -104.4887
23.5987 | -104.4876
23.6024 | -104.4867
23.606 | -104.4849
23.6097 | -104.4842
23.6133 | -104.4854
23.617 | -104.4874
23.6207 | -104.4879
23.6243 | -104.4873
23.628 | -104.4911
23.6316 | -104.4953
23.6353 | -104.4989
23.6389 | -104.5041
23.6426 | -104.5089
23.6463 | -104.5138
23.6499 | -104.5203
23.6536 | -104.5285
23.6572 | -104.5348
23.6609 | -104.5443
23.6645 | -104.5551
23.6682 | -104.5627
23.6719 | -104.5709
23.6755 | -104.5816
23.6792 | -104.5926
23.6828 | -104.602
23.6865 | -104.6142
23.6902 | -104.6251
23.6938 | -104.6355
23.6975 | -104.6476
23.7011 | -104.6587
23.7048 | -104.6708
23.7084 | -104.6828
23.7121 | -104.6936
23.7158 | -104.7073
23.7194 | -104.7191
23.7231 | -104.729
23.7267 | -104.7412
23.7304 | -104.7514
23.7341 | -104.7603
23.7377 | -104.7689
23.7414 | -104.7743
23.745 | -104.7812
23.7487 | -104.7836
23.7523 | -104.7902
23.756 | -104.797
23.7597 | -104.8046
23.7633 | -104.8103
23.767 | -104.8165
23.7706 | -104.8206
23.7743 | -104.8239
23.778 | -104.8272
23.7816 | -104.8307
23.7853 | -104.8346
23.7889 | -104.839
23.7926 | -104.8433
23.7962 | -104.8484
23.7999 | -104.8521
23.8036 | -104.8534
23.8072 | -104.8511
23.8109 | -104.8512
23.8145 | -104.8511
23.8182 | -104.8492
23.8218 | -104.8479
23.8255 | -104.8456
23.8292 | -104.8414
23.8328 | -104.8382
23.8365 | -104.8362
23.8401 | -104.8345
23.8438 | -104.8321
23.8475 | -104.83
23.8511 | -104.8279
23.8548 | -104.8267
23.8584 | -104.8277
23.8621 | -104.828
23.8657 | -104.8274
23.8694 | -104.8275
23.8731 | -104.8279
23.8767 | -104.8282
23.8804 | -104.8302
23.884 | -104.8315
23.8877 | -104.8316
23.8914 | -104.8328
23.895 | -104.8344
23.8987 | -104.838
23.9023 | -104.8431
23.906 | -104.8464
23.9096 | -104.8487
23.9133 | -104.8534
23.917 | -104.8588
23.9206 | -104.8633
23.9243 | -104.8688
23.9279 | -104.8757
23.9316 | -104.8837
23.9353 | -104.8921
23.9389 | -104.9009
23.9426 | -104.9098
23.9462 | -104.9177
23.9499 | -104.926
23.9535 | -104.9337
23.9572 | -104.9433
23.9609 | -104.9528
23.9645 | -104.9607
23.9682 | -104.9688
23.9718 | -104.9762
23.9755 | -104.9832
23.9791 | -104.9915
23.9828 | -105.0009
23.9865 | -105.0079
23.9901 | -105.0144
23.9938 | -105.0232
23.9974 | -105.03
24.0011 | -105.0354
24.0048 | -105.0434
24.0084 | -105.0533
24.0121 | -105.0614
24.0157 | -105.0667
24.0194 | -105.0721
24.023 | -105.0782
24.0267 | -105.0859
24.0304 | -105.0916
24.034 | -105.0963
24.0377 | -105.1026
24.0413 | -105.1102
24.045 | -105.117
24.0487 | -105.1254
24.0523 | -105.1326
24.056 | -105.1393
24.0596 | -105.1449
24.0633 | -105.1491
24.0669 | -105.1554
24.0706 | -105.1632
24.0743 | -105.1734
24.0779 | -105.1842
24.0816 | -105.1931
24.0852 | -105.202
24.0889 | -105.2091
24.0926 | -105.215
24.0962 | -105.221
24.0999 | -105.2274
24.1035 | -105.2339
24.1072 | -105.2385
24.1108 | -105.2413
24.1145 | -105.2471
24.1182 | -105.2538
24.1218 | -105.2604
24.1255 | -105.2653
24.1291 | -105.2692
24.1328 | -105.2724
24.1364 | -105.2756
24.1401 | -105.2804
24.1438 | -105.2854
24.1474 | -105.2906
24.1511 | -105.2937
24.1547 | -105.2985
24.1584 | -105.3041
24.1621 | -105.3088
24.1657 | -105.3137
24.1694 | -105.3183
24.173 | -105.3205
24.1767 | -105.3223
24.1803 | -105.323
24.184 | -105.3253
24.1877 | -105.3279
24.1913 | -105.33
24.195 | -105.329
24.1986 | -105.3285
24.2023 | -105.328
24.206 | -105.3268
24.2096 | -105.3267
24.2133 | -105.3261
24.2169 | -105.326
24.2206 | -105.3261
24.2242 | -105.3242
24.2279 | -105.3241
24.2316 | -105.3223
24.2352 | -105.3211
24.2389 | -105.317
24.2425 | -105.3113
24.2462 | -105.3077
24.2499 | -105.3053
24.2535 | -105.3017
24.2572 | -105.2984
24.2608 | -105.2928
24.2645 | -105.286
24.2681 | -105.2788
24.2718 | -105.2716
24.2755 | -105.2565
24.2791 | -105.2373
24.2828 | -105.2129
24.2864 | -105.1822
24.2901 | -105.1356
24.2937 | -105.0944
24.2974 | -105.0621
24.3011 | -105.0303
24.3047 | -105.0013
24.3084 | -104.9732
24.312 | -104.9428
24.3157 | -104.912
24.3194 | -104.8801
24.323 | -104.8462
24.3267 | -104.8105
24.3303 | -104.7738
24.334 | -104.7343
24.3376 | -104.6938
24.3413 | -104.6502
24.345 | -104.6057
24.3486 | -104.5583
24.3523 | -104.5085
24.3559 | -104.4579
24.3596 | -104.4086
24.3633 | -104.3592
24.3669 | -104.3066
24.3706 | -104.2537
24.3742 | -104.2009
24.3779 | -104.1481
24.3815 | -104.0953
24.3852 | -104.0431
24.3889 | -103.9914
24.3925 | -103.9398
24.3962 | -103.8894
24.3998 | -103.8408
24.4035 | -103.7945
24.4072 | -103.7494
24.4108 | -103.7044
24.4145 | -103.6612
24.4181 | -103.6192
24.4218 | -103.5812
24.4254 | -103.5471
24.4291 | -103.515
24.4328 | -103.4855
24.4364 | -103.4588
24.4401 | -103.4348
24.4437 | -103.4155
24.4474 | -103.4001
24.451 | -103.3881
24.4547 | -103.3802
24.4584 | -103.3737
24.462 | -103.3719
24.4657 | -103.3741
24.4693 | -103.3797
24.473 | -103.3893
24.4767 | -103.4025
24.4803 | -103.4188
24.484 | -103.4385
24.4876 | -103.4616
24.4913 | -103.4872
24.4949 | -103.5159
24.4986 | -103.5477
24.5023 | -103.5826
24.5059 | -103.6183
24.5096 | -103.6571
24.5132 | -103.6986
24.5169 | -103.7424
24.5206 | -103.787
24.5242 | -103.833
24.5279 | -103.88
24.5315 | -103.9289
24.5352 | -103.9775
24.5388 | -104.028
24.5425 | -104.0773
24.5462 | -104.1268
24.5498 | -104.1773
24.5535 | -104.2265
24.5571 | -104.2722
24.5608 | -104.3167
24.5645 | -104.3572
24.5681 | -104.3926
24.5718 | -104.4096
24.5754 | -104.4067
24.5791 | -104.3813
24.5827 | -104.3313
24.5864 | -104.2604
24.5901 | -104.1379
24.5937 | -104.0009
24.5974 | -103.8942
24.601 | -103.8142
24.6047 | -103.738
24.6084 | -103.6549
24.612 | -103.562
24.6157 | -103.4568
24.6193 | -103.341
24.623 | -103.2156
24.6266 | -103.0897
24.6303 | -102.9637
24.634 | -102.8367
24.6376 | -102.7109
24.6413 | -102.5727
24.6449 | -102.4235
24.6486 | -102.2703
24.6522 | -102.1135
24.6559 | -101.9534
24.6596 | -101.7933
24.6632 | -101.6333
24.6669 | -101.4753
24.6705 | -101.3189
24.6742 | -101.1638
24.6779 | -101.0116
24.6815 | -100.8631
24.6852 | -100.7175
24.6888 | -100.5767
24.6925 | -100.4403
24.6961 | -100.3087
24.6998 | -100.1821
24.7035 | -100.0611
24.7071 | -99.9458
24.7108 | -99.8361
24.7144 | -99.7323
24.7181 | -99.6353
24.7218 | -99.5453
24.7254 | -99.4625
24.7291 | -99.3864
24.7327 | -99.3175
24.7364 | -99.2562
24.74 | -99.2027
24.7437 | -99.157
24.7474 | -99.1189
24.751 | -99.0883
24.7547 | -99.0653
24.7583 | -99.0508
24.762 | -99.0435
24.7657 | -99.044
24.7693 | -99.0524
24.773 | -99.0684
24.7766 | -99.0924
24.7803 | -99.1245
24.7839 | -99.1644
24.7876 | -99.2121
24.7913 | -99.2672
24.7949 | -99.3295
24.7986 | -99.3991
24.8022 | -99.5223
24.8059 | -99.6026
24.8095 | -99.6894
24.8132 | -99.7819
24.8169 | -99.8801
24.8205 | -99.9843
24.8242 | -100.0938
24.8278 | -100.2093
24.8315 | -100.3303
24.8352 | -100.4563
24.8388 | -100.5864
24.8425 | -100.7207
24.8461 | -100.8587
24.8498 | -101.0006
24.8534 | -101.1449
24.8571 | -101.2929
24.8608 | -101.4423
24.8644 | -101.5934
24.8681 | -101.7457
24.8717 | -101.8992
24.8754 | -102.0532
24.8791 | -102.2063
24.8827 | -102.3581
24.8864 | -102.5089
24.89 | -102.6574
24.8937 | -102.8031
24.8973 | -102.9474
24.901 | -103.087
24.9047 | -103.2224
24.9083 | -103.3559
24.912 | -103.4852
24.9156 | -103.6107
24.9193 | -103.7346
24.923 | -103.8629
24.9266 | -103.9964
24.9303 | -104.1345
24.9339 | -104.2774
24.9376 | -104.4206
24.9412 | -104.5901
24.9449 | -104.7465
24.9486 | -104.8422
24.9522 | -104.8862
24.9559 | -104.9118
24.9595 | -104.9337
24.9632 | -104.9548
24.9668 | -104.9783
24.9705 | -105.0022
24.9742 | -105.0304
24.9778 | -105.0477
24.9815 | -105.056
24.9851 | -105.058
24.9888 | -105.0581
24.9925 | -105.0585
24.9961 | -105.0585
24.9998 | -105.0577
25.0034 | -105.0574
25.0071 | -105.0594
25.0107 | -105.0598
25.0144 | -105.0598
25.0181 | -105.0575
25.0217 | -105.0556
25.0254 | -105.0537
25.029 | -105.0518
25.0327 | -105.05
25.0364 | -105.0492
25.04 | -105.0465
25.0437 | -105.0443
25.0473 | -105.0412
25.051 | -105.0369
25.0546 | -105.0332
25.0583 | -105.0281
25.062 | -105.0255
25.0656 | -105.0234
25.0693 | -105.019
25.0729 | -105.0159
25.0766 | -105.0107
25.0803 | -105.0063
25.0839 | -105.0008
25.0876 | -104.9941
25.0912 | -104.9873
25.0949 | -104.9805
25.0985 | -104.9739
25.1022 | -104.9686
25.1059 | -104.9633
25.1095 | -104.9579
25.1132 | -104.9532
25.1168 | -104.9488
25.1205 | -104.9435
25.1241 | -104.9383
25.1278 | -104.9334
25.1315 | -104.9296
25.1351 | -104.9266
25.1388 | -104.9222
25.1424 | -104.9156
25.1461 | -104.911
25.1498 | -104.908
25.1534 | -104.9045
25.1571 | -104.9004
25.1607 | -104.8976
25.1644 | -104.8955
25.168 | -104.8923
25.1717 | -104.8911
25.1754 | -104.8896
25.179 | -104.8885
25.1827 | -104.8872
25.1863 | -104.8855
25.19 | -104.8833
25.1937 | -104.8806
25.1973 | -104.8789
25.201 | -104.8783
25.2046 | -104.8751
25.2083 | -104.8728
25.2119 | -104.8709
25.2156 | -104.8708
25.2193 | -104.8711
25.2229 | -104.8724
25.2266 | -104.8727
25.2302 | -104.8733
25.2339 | -104.8744
25.2376 | -104.876
25.2412 | -104.8779
25.2449 | -104.8795
25.2485 | -104.8821
25.2522 | -104.8822
25.2558 | -104.8851
25.2595 | -104.8898
25.2632 | -104.8922
25.2668 | -104.8934
25.2705 | -104.8965
25.2741 | -104.8999
25.2778 | -104.9023
25.2814 | -104.9052
25.2851 | -104.907
25.2888 | -104.9079
25.2924 | -104.9075
25.2961 | -104.9091
25.2997 | -104.9096
25.3034 | -104.9077
25.3071 | -104.9058
25.3107 | -104.9065
25.3144 | -104.9063
25.318 | -104.9051
25.3217 | -104.9022
25.3253 | -104.9008
25.329 | -104.894
25.3327 | -104.8912
25.3363 | -104.8889
25.34 | -104.8858
25.3436 | -104.8839
25.3473 | -104.8794
25.351 | -104.8737
25.3546 | -104.87
25.3583 | -104.8659
25.3619 | -104.8593
25.3656 | -104.8536
25.3692 | -104.849
25.3729 | -104.8435
25.3766 | -104.8376
25.3802 | -104.8302
25.3839 | -104.8245
25.3875 | -104.8203
25.3912 | -104.8137
25.3949 | -104.8057
25.3985 | -104.7989
25.4022 | -104.7924
25.4058 | -104.7854
25.4095 | -104.7797
25.4131 | -104.7729
25.4168 | -104.7665
25.4205 | -104.7586
25.4241 | -104.752
25.4278 | -104.7459
25.4314 | -104.7388
25.4351 | -104.731
25.4387 | -104.7241
25.4424 | -104.7176
25.4461 | -104.7107
25.4497 | -104.7036
25.4534 | -104.6986
25.457 | -104.6923
25.4607 | -104.6864
25.4644 | -104.6818
25.468 | -104.6762
25.4717 | -104.6703
25.4753 | -104.6655
25.479 | -104.6614
25.4826 | -104.6571
25.4863 | -104.6534
25.49 | -104.6488
25.4936 | -104.6431
25.4973 | -104.6383
25.5009 | -104.6347
25.5046 | -104.63
25.5083 | -104.6263
25.5119 | -104.6232
25.5156 | -104.6203
25.5192 | -104.6171
25.5229 | -104.6161
25.5265 | -104.6165
25.5302 | -104.617
25.5339 | -104.6181
25.5375 | -104.6191
25.5412 | -104.6189
25.5448 | -104.6187
25.5485 | -104.62
25.5522 | -104.6214
25.5558 | -104.6231
25.5595 | -104.6245
25.5631 | -104.627
25.5668 | -104.6293
25.5704 | -104.6315
25.5741 | -104.6331
25.5778 | -104.6352
25.5814 | -104.6379
25.5851 | -104.6405
25.5887 | -104.6433
25.5924 | -104.6478
25.596 | -104.6511
25.5997 | -104.6558
25.6034 | -104.6609
25.607 | -104.6645
25.6107 | -104.6699
25.6143 | -104.6736
25.618 | -104.6773
25.6217 | -104.6833
25.6253 | -104.6889
25.629 | -104.6934
25.6326 | -104.6982
25.6363 | -104.7045
25.6399 | -104.7098
25.6436 | -104.7156
25.6473 | -104.7216
25.6509 | -104.7281
25.6546 | -104.7337
25.6582 | -104.7389
25.6619 | -104.7452
25.6656 | -104.7515
25.6692 | -104.7563
25.6729 | -104.7613
25.6765 | -104.766
25.6802 | -104.7723
25.6838 | -104.7786
25.6875 | -104.7847
25.6912 | -104.7882
25.6948 | -104.7928
25.6985 | -104.7973
25.7021 | -104.8006
25.7058 | -104.8029
25.7095 | -104.807
25.7131 | -104.813
25.7168 | -104.8165
25.7204 | -104.8205
25.7241 | -104.8268
25.7277 | -104.8303
25.7314 | -104.8325
25.7351 | -104.8353
25.7387 | -104.84
25.7424 | -104.8445
25.746 | -104.8466
25.7497 | -104.845
25.7533 | -104.8465
25.757 | -104.8484
25.7607 | -104.8498
25.7643 | -104.8495
25.768 | -104.8495
25.7716 | -104.8488
25.7753 | -104.8477
25.779 | -104.8484
25.7826 | -104.8491
25.7863 | -104.8487
25.7899 | -104.8484
25.7936 | -104.8489
25.7972 | -104.85
25.8009 | -104.8493
25.8046 | -104.8482
25.8082 | -104.8482
25.8119 | -104.8485
25.8155 | -104.8474
25.8192 | -104.8482
25.8229 | -104.8473
25.8265 | -104.8442
25.8302 | -104.8421
25.8338 | -104.8418
25.8375 | -104.8389
25.8411 | -104.8376
25.8448 | -104.8373
25.8485 | -104.836
25.8521 | -104.8344
25.8558 | -104.8327
25.8594 | -104.8305
25.8631 | -104.8319
25.8668 | -104.8328
25.8704 | -104.8334
25.8741 | -104.8326
25.8777 | -104.8328
25.8814 | -104.8347
25.885 | -104.8357
25.8887 | -104.8348
25.8924 | -104.8351
25.896 | -104.8345
25.8997 | -104.8341
25.9033 | -104.835
25.907 | -104.8357
25.9107 | -104.8375
25.9143 | -104.8391
25.918 | -104.8424
25.9216 | -104.8447
25.9253 | -104.8446
25.9289 | -104.8473
25.9326 | -104.8495
25.9363 | -104.8502
25.9399 | -104.8519
25.9436 | -104.8535
25.9472 | -104.8548
25.9509 | -104.8563
25.9545 | -104.8577
25.9582 | -104.8595
25.9619 | -104.8613
25.9655 | -104.863
25.9692 | -104.8651
25.9728 | -104.8669
25.9765 | -104.8695
25.9802 | -104.8718
25.9838 | -104.873
25.9875 | -104.8743
25.9911 | -104.877
25.9948 | -104.8787
25.9984 | -104.8808
26.0021 | -104.8842
26.0058 | -104.8873
26.0094 | -104.8909
26.0131 | -104.8956
26.0167 | -104.8991
26.0204 | -104.9015
26.0241 | -104.9046
26.0277 | -104.9057
26.0314 | -104.9059
26.035 | -104.9067
26.0387 | -104.9088
26.0423 | -104.9129
26.046 | -104.9173
26.0497 | -104.9194
26.0533 | -104.9223
26.057 | -104.9254
26.0606 | -104.9274
26.0643 | -104.9296
26.0679 | -104.9308
26.0716 | -104.9326
26.0753 | -104.9357
26.0789 | -104.9362
26.0826 | -104.9363
26.0862 | -104.9365
26.0899 | -104.9358
26.0936 | -104.9366
26.0972 | -104.9385
26.1009 | -104.9397
26.1045 | -104.9376
26.1082 | -104.936
26.1118 | -104.9357
26.1155 | -104.9377
26.1192 | -104.9374
26.1228 | -104.9349
26.1265 | -104.9319
26.1301 | -104.9311
26.1338 | -104.9308
26.1375 | -104.9292
26.1411 | -104.9263
26.1448 | -104.9195
26.1484 | -104.9132
26.1521 | -104.9083
26.1557 | -104.9039
26.1594 | -104.898
26.1631 | -104.8928
26.1667 | -104.8886
26.1704 | -104.8829
26.174 | -104.8776
26.1777 | -104.8725
26.1814 | -104.8659
26.185 | -104.8592
26.1887 | -104.8534
26.1923 | -104.8485
26.196 | -104.8444
26.1996 | -104.8403
26.2033 | -104.8372
26.207 | -104.8326
26.2106 | -104.8279
26.2143 | -104.8235
26.2179 | -104.8209
26.2216 | -104.8203
26.2252 | -104.8184
26.2289 | -104.8161
26.2326 | -104.8124
26.2362 | -104.8102
26.2399 | -104.8077
26.2435 | -104.8059
26.2472 | -104.805
26.2509 | -104.8035
26.2545 | -104.8028
26.2582 | -104.802
26.2618 | -104.801
26.2655 | -104.8022
26.2691 | -104.8048
26.2728 | -104.807
26.2765 | -104.809
26.2801 | -104.8111
26.2838 | -104.8138
26.2874 | -104.8184
26.2911 | -104.8246
26.2948 | -104.8325
26.2984 | -104.8377
26.3021 | -104.8442
26.3057 | -104.8521
26.3094 | -104.8588
26.313 | -104.8665
26.3167 | -104.8763
26.3204 | -104.8866
26.324 | -104.896
26.3277 | -104.9056
26.3313 | -104.9188
26.335 | -104.9313
26.3387 | -104.9445
26.3423 | -104.9586
26.346 | -104.9715
26.3496 | -104.9849
26.3533 | -104.9991
26.3569 | -105.0128
26.3606 | -105.0263
26.3643 | -105.0406
26.3679 | -105.0539
26.3716 | -105.0658
26.3752 | -105.0778
26.3789 | -105.0893
26.3826 | -105.1018
26.3862 | -105.1085
26.3899 | -105.1188
26.3935 | -105.1297
26.3972 | -105.1407
26.4008 | -105.1527
26.4045 | -105.1639
26.4082 | -105.1738
26.4118 | -105.1821
26.4155 | -105.1891
26.4191 | -105.1941
26.4228 | -105.1999
26.4264 | -105.2054
26.4301 | -105.2119
26.4338 | -105.2171
26.4374 | -105.2196
26.4411 | -105.2235
26.4447 | -105.2263
26.4484 | -105.2263
26.4521 | -105.2276
26.4557 | -105.229
26.4594 | -105.2295
26.463 | -105.2297
26.4667 | -105.2295
26.4703 | -105.2288
26.474 | -105.2275
26.4777 | -105.2274
26.4813 | -105.2267
26.485 | -105.2264
26.4886 | -105.225
26.4923 | -105.2231
26.496 | -105.2229
26.4996 | -105.2235
26.5033 | -105.2206
26.5069 | -105.2171
26.5106 | -105.2151
26.5142 | -105.215
26.5179 | -105.2167
26.5216 | -105.2185
26.5252 | -105.2187
26.5289 | -105.218
26.5325 | -105.2179
26.5362 | -105.2167
26.5399 | -105.2165
26.5435 | -105.2177
26.5472 | -105.2179
26.5508 | -105.2189
26.5545 | -105.2206
26.5581 | -105.2214
26.5618 | -105.2233
26.5655 | -105.2261
26.5691 | -105.2267
26.5728 | -105.2271
26.5764 | -105.2285
26.5801 | -105.2306
26.5837 | -105.2327
26.5874 | -105.2355
26.5911 | -105.2372
26.5947 | -105.2393
26.5984 | -105.2419
26.602 | -105.2444
26.6057 | -105.2473
26.6094 | -105.2478
26.613 | -105.2489
26.6167 | -105.2518
26.6203 | -105.253
26.624 | -105.2545
26.6276 | -105.2583
26.6313 | -105.2616
26.635 | -105.2638
26.6386 | -105.2651
26.6423 | -105.2669
26.6459 | -105.2688
26.6496 | -105.2716
26.6533 | -105.2729
26.6569 | -105.2747
26.6606 | -105.2738
26.6642 | -105.2725
26.6679 | -105.2724
26.6715 | -105.2729
26.6752 | -105.2736
26.6789 | -105.2739
26.6825 | -105.2747
26.6862 | -105.2741
26.6898 | -105.2715
26.6935 | -105.27
26.6972 | -105.2687
26.7008 | -105.2666
26.7045 | -105.262
26.7081 | -105.2595
26.7118 | -105.2573
26.7154 | -105.2562
26.7191 | -105.2538
26.7228 | -105.2518
26.7264 | -105.2503
26.7301 | -105.2483
26.7337 | -105.2456
26.7374 | -105.2442
26.741 | -105.2425
26.7447 | -105.24
26.7484 | -105.2398
26.752 | -105.2395
26.7557 | -105.2365
26.7593 | -105.2349
26.763 | -105.233
26.7667 | -105.2311
26.7703 | -105.2301
26.774 | -105.2282
26.7776 | -105.225
26.7813 | -105.2235
26.7849 | -105.2218
26.7886 | -105.219
26.7923 | -105.2165
26.7959 | -105.2158
26.7996 | -105.2153
26.8032 | -105.2126
26.8069 | -105.2092
26.8106 | -105.207
26.8142 | -105.206
26.8179 | -105.2054
26.8215 | -105.2039
26.8252 | -105.2034
26.8288 | -105.2029
26.8325 | -105.2014
26.8362 | -105.1998
26.8398 | -105.1991
26.8435 | -105.1982
26.8471 | -105.1962
26.8508 | -105.1943
26.8545 | -105.1919
26.8581 | -105.1894
26.8618 | -105.1856
26.8654 | -105.1825
26.8691 | -105.1809
26.8727 | -105.1789
26.8764 | -105.1748
26.8801 | -105.1729
26.8837 | -105.1703
26.8874 | -105.1666
26.891 | -105.1624
26.8947 | -105.1577
26.8983 | -105.1529
26.902 | -105.1477
26.9057 | -105.1426
26.9093 | -105.1387
26.913 | -105.1329
26.9166 | -105.128
26.9203 | -105.1226
26.924 | -105.1173
26.9276 | -105.1116
26.9313 | -105.105
26.9349 | -105.0967
26.9386 | -105.0904
26.9422 | -105.0826
26.9459 | -105.075
26.9496 | -105.0684
26.9532 | -105.0599
26.9569 | -105.0506
26.9605 | -105.0397
26.9642 | -105.03
26.9679 | -105.0208
26.9715 | -105.0112
26.9752 | -105.0008
26.9788 | -104.9903
26.9825 | -104.9807
26.9861 | -104.972
26.9898 | -104.9648
26.9935 | -104.9572
26.9971 | -104.9471
27.0008 | -104.9382
27.0044 | -104.9283
27.0081 | -104.9157
27.0118 | -104.903
27.0154 | -104.8921
27.0191 | -104.8791
27.0227 | -104.8668
27.0264 | -104.8552
27.03 | -104.8446
27.0337 | -104.8341
27.0374 | -104.8207
27.041 | -104.8098
27.0447 | -104.7977
27.0483 | -104.787
27.052 | -104.7769
27.0556 | -104.7642
27.0593 | -104.751
27.063 | -104.7389
27.0666 | -104.7284
27.0703 | -104.7175
27.0739 | -104.7064
27.0776 | -104.6969
27.0813 | -104.687
27.0849 | -104.6767
27.0886 | -104.6652
27.0922 | -104.6534
27.0959 | -104.6415
27.0995 | -104.6299
27.1032 | -104.6188
27.1069 | -104.6072
27.1105 | -104.5971
27.1142 | -104.5886
27.1178 | -104.5781
27.1215 | -104.5671
27.1252 | -104.5589
27.1288 | -104.5498
27.1325 | -104.5408
27.1361 | -104.5325
27.1398 | -104.5254
27.1434 | -104.5198
27.1471 | -104.5137
27.1508 | -104.5072
27.1544 | -104.5018
27.1581 | -104.4984
27.1617 | -104.4948
27.1654 | -104.4921
27.1691 | -104.4902
27.1727 | -104.4886
27.1764 | -104.4861
27.18 | -104.4828
27.1837 | -104.4815
27.1873 | -104.4812
27.191 | -104.4822
27.1947 | -104.4844
27.1983 | -104.486
27.202 | -104.4868
27.2056 | -104.4883
27.2093 | -104.4893
27.2129 | -104.4909
27.2166 | -104.4934
27.2203 | -104.4968
27.2239 | -104.5012
27.2276 | -104.5068
27.2312 | -104.5127
27.2349 | -104.5185
27.2386 | -104.5236
27.2422 | -104.5295
27.2459 | -104.5358
27.2495 | -104.5424
27.2532 | -104.5495
27.2568 | -104.5568
27.2605 | -104.5634
27.2642 | -104.5699
27.2678 | -104.5768
27.2715 | -104.5834
27.2751 | -104.5903
27.2788 | -104.5967
27.2825 | -104.6042
27.2861 | -104.6121
27.2898 | -104.6184
27.2934 | -104.6243
27.2971 | -104.6318
27.3007 | -104.6402
27.3044 | -104.6474
27.3081 | -104.6546
27.3117 | -104.6622
27.3154 | -104.6688
27.319 | -104.676
27.3227 | -104.6812
27.3264 | -104.6889
27.33 | -104.6968
27.3337 | -104.7033
27.3373 | -104.7109
27.341 | -104.7172
27.3446 | -104.7246
27.3483 | -104.732
27.352 | -104.7383
27.3556 | -104.7442
27.3593 | -104.7492
27.3629 | -104.7535
27.3666 | -104.7577
27.3702 | -104.7621
27.3739 | -104.7653
27.3776 | -104.7693
27.3812 | -104.774
27.3849 | -104.7766
27.3885 | -104.7797
27.3922 | -104.7849
27.3959 | -104.7906
27.3995 | -104.7948
27.4032 | -104.8007
27.4068 | -104.8042
27.4105 | -104.8062
27.4141 | -104.8073
27.4178 | -104.8091
27.4215 | -104.8122
27.4251 | -104.8144
27.4288 | -104.818
27.4324 | -104.8197
27.4361 | -104.8204
27.4398 | -104.8225
27.4434 | -104.8245
27.4471 | -104.8272
27.4507 | -104.8292
27.4544 | -104.8301
27.458 | -104.8311
27.4617 | -104.8319
27.4654 | -104.833
27.469 | -104.8337
27.4727 | -104.8358
27.4763 | -104.8377
27.48 | -104.8394
27.4837 | -104.8419
27.4873 | -104.8445
27.491 | -104.8469
27.4946 | -104.85
27.4983 | -104.8503
27.5019 | -104.8491
27.5056 | -104.8495
27.5093 | -104.8515
27.5129 | -104.8518
27.5166 | -104.8519
27.5202 | -104.8523
27.5239 | -104.8535
27.5276 | -104.8544
27.5312 | -104.8547
27.5349 | -104.8564
27.5385 | -104.8563
27.5422 | -104.8567
27.5458 | -104.8576
27.5495 | -104.8588
27.5532 | -104.8579
27.5568 | -104.8578
27.5605 | -104.8577
27.5641 | -104.8589
27.5678 | -104.8625
27.5714 | -104.8649
27.5751 | -104.8641
27.5788 | -104.8622
27.5824 | -104.8603
27.5861 | -104.8601
27.5897 | -104.8608
27.5934 | -104.8617
27.5971 | -104.8631
27.6007 | -104.8639
27.6044 | -104.8651
27.608 | -104.8654
27.6117 | -104.8643
27.6153 | -104.864
27.619 | -104.8644
27.6227 | -104.8643
27.6263 | -104.8638
27.63 | -104.8628
27.6336 | -104.8603
27.6373 | -104.8572
27.641 | -104.8536
27.6446 | -104.8505
27.6483 | -104.8478
27.6519 | -104.8469
27.6556 | -104.8474
27.6592 | -104.8462
27.6629 | -104.8458
27.6666 | -104.8461
27.6702 | -104.8446
27.6739 | -104.8422
27.6775 | -104.8405
27.6812 | -104.838
27.6849 | -104.8355
27.6885 | -104.8331
27.6922 | -104.8319
27.6958 | -104.8302
27.6995 | -104.8275
27.7031 | -104.8233
27.7068 | -104.8208
27.7105 | -104.8191
27.7141 | -104.8158
27.7178 | -104.8114
27.7214 | -104.8077
27.7251 | -104.8054
27.7287 | -104.8026
27.7324 | -104.7975
27.7361 | -104.792
27.7397 | -104.7878
27.7434 | -104.7841
27.747 | -104.7811
27.7507 | -104.7774
27.7544 | -104.7738
27.758 | -104.7684
27.7617 | -104.7639
27.7653 | -104.758
27.769 | -104.7526
27.7726 | -104.7488
27.7763 | -104.7429
27.78 | -104.7366
27.7836 | -104.7308
27.7873 | -104.7242
27.7909 | -104.7172
27.7946 | -104.7118
27.7983 | -104.7062
27.8019 | -104.7009
27.8056 | -104.6959
27.8092 | -104.6917
27.8129 | -104.686
27.8165 | -104.6808
27.8202 | -104.676
27.8239 | -104.6723
27.8275 | -104.6689
27.8312 | -104.665
27.8348 | -104.6633
27.8385 | -104.6624
27.8422 | -104.661
27.8458 | -104.6594
27.8495 | -104.659
27.8531 | -104.6579
27.8568 | -104.6552
27.8604 | -104.6538
27.8641 | -104.6518
27.8678 | -104.651
27.8714 | -104.6523
27.8751 | -104.6539
27.8787 | -104.656
27.8824 | -104.6574
27.886 | -104.6579
27.8897 | -104.6613
27.8934 | -104.6659
27.897 | -104.6703
27.9007 | -104.6736
27.9043 | -104.6775
27.908 | -104.6823
27.9117 | -104.6855
27.9153 | -104.69
27.919 | -104.6935
27.9226 | -104.6977
27.9263 | -104.6996
27.9299 | -104.7032
27.9336 | -104.7072
27.9373 | -104.7095
27.9409 | -104.7116
27.9446 | -104.7157
27.9482 | -104.7206
27.9519 | -104.7255
27.9556 | -104.7304
27.9592 | -104.7341
27.9629 | -104.7377
27.9665 | -104.7422
27.9702 | -104.7428
27.9738 | -104.7462
27.9775 | -104.7488
27.9812 | -104.7508
27.9848 | -104.7537
27.9885 | -104.7556
27.9921 | -104.7564
27.9958 | -104.7574
27.9995 | -104.757
28.0031 | -104.7568