## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 245 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 262
Hood | 02-92B/P | 189
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1005 | hPa |  | PRC PRO 01
Humidity | 54.9 | %/Hr
Internal temperature of calibration bench | 23.96 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1012 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.8 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.03 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.03 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3070 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 06/06/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 12/06/2018
Technical Manager | Anthony HUE | Approval |  | 13/06/2018
CEO | Robert GLEMOT | Approval |  | 13/06/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3985 | 57.5154
0.012209 | -35.9718 | 38.3371
0.018313 | -34.9068 | 28.2182
0.024417 | -34.4413 | 21.9834
0.030521 | -34.3304 | 18.2237
0.036626 | -34.1854 | 15.1648
0.04273 | -34.0894 | 12.7978
0.048834 | -34.0431 | 11.2867
0.054938 | -33.9775 | 10.1594
0.061043 | -33.9157 | 9.0771
0.067147 | -33.9049 | 8.2555
0.073251 | -33.9058 | 7.4979
0.079355 | -33.9244 | 6.8385
0.08546 | -33.9386 | 6.2445
0.091564 | -33.9336 | 5.8038
0.097668 | -33.9149 | 5.5567
0.10377 | -33.8922 | 5.3109
0.10988 | -33.8745 | 5.0342
0.11598 | -33.865 | 4.7411
0.12209 | -33.8668 | 4.4563
0.12819 | -33.8767 | 4.1718
0.13429 | -33.8877 | 3.8832
0.1404 | -33.8911 | 3.6049
0.1465 | -33.893 | 3.347
0.15261 | -33.8902 | 3.1269
0.15871 | -33.8836 | 2.9652
0.16482 | -33.8753 | 2.8634
0.17092 | -33.8728 | 2.8256
0.17702 | -33.8705 | 2.7543
0.18313 | -33.8694 | 2.6768
0.18923 | -33.8688 | 2.5682
0.19534 | -33.8674 | 2.4415
0.20144 | -33.8659 | 2.2992
0.20754 | -33.8628 | 2.1784
0.21365 | -33.8617 | 2.0638
0.21975 | -33.8611 | 1.9662
0.22586 | -33.862 | 1.8791
0.23196 | -33.8644 | 1.8072
0.23807 | -33.8685 | 1.7453
0.24417 | -33.8726 | 1.688
0.25027 | -33.8765 | 1.6267
0.25638 | -33.8789 | 1.5659
0.26248 | -33.8806 | 1.5002
0.26859 | -33.8815 | 1.4356
0.27469 | -33.8818 | 1.3649
0.2808 | -33.8827 | 1.2908
0.2869 | -33.8834 | 1.2197
0.293 | -33.8842 | 1.1523
0.29911 | -33.885 | 1.0919
0.30521 | -33.8859 | 1.0323
0.31132 | -33.8866 | 0.99339
0.31742 | -33.8877 | 0.95028
0.32353 | -33.888 | 0.91518
0.32963 | -33.8888 | 0.88618
0.33573 | -33.8898 | 0.8662
0.34184 | -33.8911 | 0.84524
0.34794 | -33.8931 | 0.82241
0.35405 | -33.8952 | 0.79906
0.36015 | -33.8976 | 0.77459
0.36626 | -33.8996 | 0.74806
0.37236 | -33.9016 | 0.71231
0.37846 | -33.9037 | 0.6689
0.38457 | -33.9044 | 0.61214
0.39067 | -33.9054 | 0.56603
0.39678 | -33.9059 | 0.51446
0.40288 | -33.9063 | 0.46061
0.40899 | -33.9061 | 0.40872
0.41509 | -33.9053 | 0.35772
0.42119 | -33.9052 | 0.31411
0.4273 | -33.9051 | 0.26764
0.4334 | -33.9049 | 0.22123
0.43951 | -33.9051 | 0.17758
0.44561 | -33.9059 | 0.13763
0.45172 | -33.9075 | 0.10577
0.45782 | -33.9082 | 0.068514
0.46392 | -33.9094 | 0.03213
0.47003 | -33.9105 | -0.00036112
0.47613 | -33.9113 | -0.034554
0.48224 | -33.9118 | -0.069491
0.48834 | -33.9108 | -0.10433
0.49445 | -33.9106 | -0.13972
0.50055 | -33.91 | -0.17183
0.50665 | -33.9092 | -0.2061
0.51276 | -33.9081 | -0.24016
0.51886 | -33.9071 | -0.2741
0.52497 | -33.9058 | -0.29815
0.53107 | -33.9049 | -0.32756
0.53717 | -33.9042 | -0.35569
0.54328 | -33.9033 | -0.37835
0.54938 | -33.903 | -0.40068
0.55549 | -33.903 | -0.4214
0.56159 | -33.9041 | -0.43278
0.5677 | -33.9043 | -0.44713
0.5738 | -33.9054 | -0.46329
0.5799 | -33.9061 | -0.47805
0.58601 | -33.9072 | -0.48802
0.59211 | -33.9088 | -0.49414
0.59822 | -33.9102 | -0.50346
0.60432 | -33.9117 | -0.51537
0.61043 | -33.9129 | -0.52318
0.61653 | -33.9141 | -0.53333
0.62263 | -33.9157 | -0.54256
0.62874 | -33.9163 | -0.5531
0.63484 | -33.9172 | -0.56224
0.64095 | -33.9177 | -0.57232
0.64705 | -33.918 | -0.58196
0.65316 | -33.9182 | -0.59428
0.65926 | -33.9183 | -0.60997
0.66536 | -33.9171 | -0.62726
0.67147 | -33.9167 | -0.64032
0.67757 | -33.9163 | -0.65447
0.68368 | -33.9155 | -0.67325
0.68978 | -33.9147 | -0.69033
0.69589 | -33.9137 | -0.70922
0.70199 | -33.9127 | -0.73196
0.70809 | -33.9117 | -0.75106
0.7142 | -33.911 | -0.77147
0.7203 | -33.9098 | -0.79064
0.72641 | -33.9091 | -0.81313
0.73251 | -33.9084 | -0.83314
0.73862 | -33.9086 | -0.85502
0.74472 | -33.9082 | -0.87507
0.75082 | -33.9079 | -0.89724
0.75693 | -33.9082 | -0.91909
0.76303 | -33.9087 | -0.94113
0.76914 | -33.9105 | -0.96167
0.77524 | -33.9114 | -0.98136
0.78135 | -33.9127 | -1.0015
0.78745 | -33.914 | -1.0215
0.79355 | -33.9152 | -1.0413
0.79966 | -33.917 | -1.0608
0.80576 | -33.9193 | -1.0804
0.81187 | -33.9209 | -1.102
0.81797 | -33.9228 | -1.1233
0.82408 | -33.9245 | -1.1438
0.83018 | -33.9262 | -1.1643
0.83628 | -33.9281 | -1.1827
0.84239 | -33.9293 | -1.2032
0.84849 | -33.9307 | -1.2248
0.8546 | -33.9321 | -1.2446
0.8607 | -33.9334 | -1.2653
0.86681 | -33.9344 | -1.2842
0.87291 | -33.9354 | -1.3029
0.87901 | -33.9358 | -1.3208
0.88512 | -33.9366 | -1.3388
0.89122 | -33.9373 | -1.357
0.89733 | -33.9376 | -1.3725
0.90343 | -33.9381 | -1.39
0.90953 | -33.9382 | -1.4042
0.91564 | -33.9382 | -1.42
0.92174 | -33.9384 | -1.4348
0.92785 | -33.9384 | -1.4493
0.93395 | -33.9385 | -1.4631
0.94006 | -33.9386 | -1.4758
0.94616 | -33.9383 | -1.4876
0.95226 | -33.9384 | -1.5002
0.95837 | -33.9385 | -1.514
0.96447 | -33.9386 | -1.5267
0.97058 | -33.9385 | -1.54
0.97668 | -33.9382 | -1.5533
0.98279 | -33.9379 | -1.5659
0.98889 | -33.9377 | -1.5765
0.99499 | -33.9376 | -1.5898
1.0011 | -33.9375 | -1.6004
1.0072 | -33.9371 | -1.6119
1.0133 | -33.9367 | -1.6225
1.0194 | -33.9363 | -1.6354
1.0255 | -33.9359 | -1.649
1.0316 | -33.9353 | -1.6607
1.0377 | -33.9349 | -1.6743
1.0438 | -33.9343 | -1.6869
1.0499 | -33.9335 | -1.702
1.056 | -33.9327 | -1.7162
1.0621 | -33.9321 | -1.7303
1.0682 | -33.9316 | -1.7426
1.0743 | -33.9309 | -1.7578
1.0805 | -33.9303 | -1.7721
1.0866 | -33.9293 | -1.7869
1.0927 | -33.9285 | -1.8013
1.0988 | -33.9278 | -1.816
1.1049 | -33.9269 | -1.83
1.111 | -33.9261 | -1.8452
1.1171 | -33.9253 | -1.8596
1.1232 | -33.9248 | -1.8745
1.1293 | -33.9244 | -1.8884
1.1354 | -33.924 | -1.9032
1.1415 | -33.9236 | -1.9185
1.1476 | -33.9229 | -1.9338
1.1537 | -33.9224 | -1.9483
1.1598 | -33.9225 | -1.9628
1.1659 | -33.9222 | -1.9759
1.172 | -33.9219 | -1.9896
1.1781 | -33.9216 | -2.0033
1.1842 | -33.9214 | -2.0173
1.1903 | -33.9218 | -2.0281
1.1964 | -33.9216 | -2.0413
1.2025 | -33.9217 | -2.055
1.2086 | -33.922 | -2.0672
1.2147 | -33.922 | -2.0795
1.2209 | -33.9224 | -2.0905
1.227 | -33.9231 | -2.101
1.2331 | -33.9235 | -2.1128
1.2392 | -33.9238 | -2.1245
1.2453 | -33.9242 | -2.1353
1.2514 | -33.9246 | -2.145
1.2575 | -33.9252 | -2.1558
1.2636 | -33.926 | -2.1646
1.2697 | -33.9266 | -2.1736
1.2758 | -33.9273 | -2.1846
1.2819 | -33.928 | -2.1944
1.288 | -33.9287 | -2.205
1.2941 | -33.9295 | -2.2132
1.3002 | -33.9302 | -2.2211
1.3063 | -33.9307 | -2.23
1.3124 | -33.9314 | -2.2388
1.3185 | -33.9321 | -2.2477
1.3246 | -33.9328 | -2.2566
1.3307 | -33.9336 | -2.265
1.3368 | -33.9344 | -2.2735
1.3429 | -33.9351 | -2.2823
1.349 | -33.9356 | -2.2914
1.3551 | -33.9363 | -2.3006
1.3613 | -33.9368 | -2.3087
1.3674 | -33.9374 | -2.3185
1.3735 | -33.9379 | -2.3281
1.3796 | -33.9383 | -2.337
1.3857 | -33.9386 | -2.346
1.3918 | -33.9392 | -2.356
1.3979 | -33.9397 | -2.3649
1.404 | -33.9402 | -2.376
1.4101 | -33.9406 | -2.3862
1.4162 | -33.941 | -2.3966
1.4223 | -33.9415 | -2.4077
1.4284 | -33.9419 | -2.4175
1.4345 | -33.9422 | -2.4288
1.4406 | -33.9425 | -2.4421
1.4467 | -33.9429 | -2.4535
1.4528 | -33.9434 | -2.4646
1.4589 | -33.9438 | -2.4763
1.465 | -33.9441 | -2.4869
1.4711 | -33.9445 | -2.498
1.4772 | -33.945 | -2.5111
1.4833 | -33.9455 | -2.524
1.4894 | -33.946 | -2.5359
1.4955 | -33.9465 | -2.5477
1.5016 | -33.9471 | -2.56
1.5078 | -33.9474 | -2.5736
1.5139 | -33.9478 | -2.5865
1.52 | -33.9482 | -2.5988
1.5261 | -33.9485 | -2.6113
1.5322 | -33.9489 | -2.6251
1.5383 | -33.9492 | -2.6374
1.5444 | -33.9498 | -2.6499
1.5505 | -33.9502 | -2.6631
1.5566 | -33.9507 | -2.6759
1.5627 | -33.9512 | -2.6894
1.5688 | -33.9518 | -2.7028
1.5749 | -33.9521 | -2.7152
1.581 | -33.9524 | -2.7283
1.5871 | -33.9527 | -2.7418
1.5932 | -33.9531 | -2.7545
1.5993 | -33.9535 | -2.7685
1.6054 | -33.9538 | -2.7815
1.6115 | -33.9542 | -2.7954
1.6176 | -33.9545 | -2.8078
1.6237 | -33.9547 | -2.8212
1.6298 | -33.9549 | -2.8335
1.6359 | -33.9553 | -2.8461
1.642 | -33.9557 | -2.8577
1.6482 | -33.9559 | -2.8705
1.6543 | -33.9562 | -2.884
1.6604 | -33.9565 | -2.8958
1.6665 | -33.9565 | -2.908
1.6726 | -33.9567 | -2.9203
1.6787 | -33.957 | -2.9326
1.6848 | -33.9569 | -2.945
1.6909 | -33.9572 | -2.9575
1.697 | -33.9573 | -2.9695
1.7031 | -33.9574 | -2.9818
1.7092 | -33.9573 | -2.9937
1.7153 | -33.9574 | -3.0055
1.7214 | -33.9572 | -3.0175
1.7275 | -33.9571 | -3.0299
1.7336 | -33.957 | -3.0426
1.7397 | -33.9571 | -3.0548
1.7458 | -33.9571 | -3.0667
1.7519 | -33.9571 | -3.0787
1.758 | -33.9571 | -3.0905
1.7641 | -33.9572 | -3.1019
1.7702 | -33.9571 | -3.1146
1.7763 | -33.9571 | -3.1259
1.7824 | -33.957 | -3.1382
1.7885 | -33.957 | -3.1505
1.7947 | -33.957 | -3.1623
1.8008 | -33.957 | -3.1734
1.8069 | -33.957 | -3.1847
1.813 | -33.957 | -3.1968
1.8191 | -33.957 | -3.2093
1.8252 | -33.957 | -3.2219
1.8313 | -33.9571 | -3.234
1.8374 | -33.9572 | -3.2457
1.8435 | -33.9573 | -3.2575
1.8496 | -33.9573 | -3.2704
1.8557 | -33.9575 | -3.283
1.8618 | -33.9577 | -3.2951
1.8679 | -33.9578 | -3.307
1.874 | -33.9579 | -3.3189
1.8801 | -33.958 | -3.3312
1.8862 | -33.9582 | -3.343
1.8923 | -33.9583 | -3.3551
1.8984 | -33.9584 | -3.3671
1.9045 | -33.9585 | -3.3785
1.9106 | -33.9585 | -3.3907
1.9167 | -33.9585 | -3.4027
1.9228 | -33.9586 | -3.4145
1.9289 | -33.9587 | -3.4263
1.9351 | -33.9589 | -3.437
1.9412 | -33.9591 | -3.4482
1.9473 | -33.9593 | -3.4602
1.9534 | -33.9595 | -3.4715
1.9595 | -33.9597 | -3.4832
1.9656 | -33.9599 | -3.4955
1.9717 | -33.9601 | -3.5071
1.9778 | -33.9604 | -3.5184
1.9839 | -33.9607 | -3.5298
1.99 | -33.9608 | -3.5408
1.9961 | -33.9609 | -3.5529
2.0022 | -33.9611 | -3.5642
2.0083 | -33.9612 | -3.5754
2.0144 | -33.9614 | -3.5867
2.0205 | -33.9617 | -3.5984
2.0266 | -33.962 | -3.6097
2.0327 | -33.9622 | -3.6212
2.0388 | -33.9626 | -3.6327
2.0449 | -33.9628 | -3.644
2.051 | -33.963 | -3.6553
2.0571 | -33.9634 | -3.6666
2.0632 | -33.9635 | -3.678
2.0693 | -33.9637 | -3.6901
2.0754 | -33.964 | -3.7013
2.0816 | -33.9643 | -3.7134
2.0877 | -33.9645 | -3.7248
2.0938 | -33.9648 | -3.7368
2.0999 | -33.965 | -3.7485
2.106 | -33.9652 | -3.76
2.1121 | -33.9653 | -3.7718
2.1182 | -33.9656 | -3.7825
2.1243 | -33.9658 | -3.7932
2.1304 | -33.966 | -3.8042
2.1365 | -33.9663 | -3.8159
2.1426 | -33.9665 | -3.8277
2.1487 | -33.9668 | -3.8388
2.1548 | -33.9672 | -3.8506
2.1609 | -33.9676 | -3.8619
2.167 | -33.968 | -3.873
2.1731 | -33.9683 | -3.8849
2.1792 | -33.9685 | -3.8958
2.1853 | -33.9689 | -3.9076
2.1914 | -33.9692 | -3.9189
2.1975 | -33.9696 | -3.9303
2.2036 | -33.9698 | -3.9423
2.2097 | -33.97 | -3.9537
2.2158 | -33.9704 | -3.9642
2.222 | -33.9706 | -3.9745
2.2281 | -33.9709 | -3.9858
2.2342 | -33.9712 | -3.997
2.2403 | -33.9714 | -4.0075
2.2464 | -33.9716 | -4.0181
2.2525 | -33.9719 | -4.0294
2.2586 | -33.9722 | -4.0399
2.2647 | -33.9725 | -4.0502
2.2708 | -33.9728 | -4.0612
2.2769 | -33.9731 | -4.0718
2.283 | -33.9733 | -4.0828
2.2891 | -33.9735 | -4.0936
2.2952 | -33.9737 | -4.104
2.3013 | -33.9739 | -4.1147
2.3074 | -33.9742 | -4.125
2.3135 | -33.9744 | -4.1363
2.3196 | -33.9746 | -4.1463
2.3257 | -33.9748 | -4.1569
2.3318 | -33.9749 | -4.1678
2.3379 | -33.9751 | -4.1785
2.344 | -33.9752 | -4.1897
2.3501 | -33.9754 | -4.2003
2.3562 | -33.9756 | -4.2105
2.3623 | -33.9757 | -4.2208
2.3685 | -33.9758 | -4.2312
2.3746 | -33.9759 | -4.2423
2.3807 | -33.9761 | -4.2526
2.3868 | -33.9763 | -4.2632
2.3929 | -33.9765 | -4.2737
2.399 | -33.9765 | -4.2841
2.4051 | -33.9768 | -4.2944
2.4112 | -33.977 | -4.3049
2.4173 | -33.9772 | -4.3154
2.4234 | -33.9772 | -4.3261
2.4295 | -33.9774 | -4.3368
2.4356 | -33.9775 | -4.3475
2.4417 | -33.9776 | -4.3586
2.4478 | -33.9777 | -4.3693
2.4539 | -33.9779 | -4.3794
2.46 | -33.978 | -4.3907
2.4661 | -33.9781 | -4.4016
2.4722 | -33.9783 | -4.4117
2.4783 | -33.9783 | -4.4219
2.4844 | -33.9784 | -4.4328
2.4905 | -33.9785 | -4.4437
2.4966 | -33.9786 | -4.4543
2.5027 | -33.9787 | -4.4647
2.5089 | -33.9788 | -4.4754
2.515 | -33.9788 | -4.4861
2.5211 | -33.9789 | -4.4967
2.5272 | -33.9789 | -4.5075
2.5333 | -33.9791 | -4.5183
2.5394 | -33.9792 | -4.5289
2.5455 | -33.9793 | -4.5402
2.5516 | -33.9794 | -4.5504
2.5577 | -33.9796 | -4.5617
2.5638 | -33.9797 | -4.5733
2.5699 | -33.9798 | -4.5846
2.576 | -33.9798 | -4.596
2.5821 | -33.9799 | -4.6071
2.5882 | -33.98 | -4.6181
2.5943 | -33.9801 | -4.6288
2.6004 | -33.9802 | -4.6396
2.6065 | -33.9803 | -4.6499
2.6126 | -33.9804 | -4.6602
2.6187 | -33.9805 | -4.6705
2.6248 | -33.9806 | -4.6812
2.6309 | -33.9808 | -4.6925
2.637 | -33.981 | -4.7039
2.6431 | -33.9811 | -4.7142
2.6492 | -33.9812 | -4.7255
2.6554 | -33.9814 | -4.7362
2.6615 | -33.9815 | -4.747
2.6676 | -33.9817 | -4.7583
2.6737 | -33.982 | -4.7693
2.6798 | -33.9822 | -4.7802
2.6859 | -33.9824 | -4.7914
2.692 | -33.9825 | -4.8021
2.6981 | -33.9827 | -4.8131
2.7042 | -33.9828 | -4.8241
2.7103 | -33.983 | -4.8346
2.7164 | -33.9833 | -4.8454
2.7225 | -33.9835 | -4.8566
2.7286 | -33.9837 | -4.8674
2.7347 | -33.9839 | -4.8778
2.7408 | -33.9842 | -4.8878
2.7469 | -33.9843 | -4.899
2.753 | -33.9846 | -4.9094
2.7591 | -33.9849 | -4.9205
2.7652 | -33.9851 | -4.9319
2.7713 | -33.9852 | -4.9424
2.7774 | -33.9855 | -4.9532
2.7835 | -33.9856 | -4.9641
2.7896 | -33.9858 | -4.9744
2.7958 | -33.9861 | -4.9851
2.8019 | -33.9863 | -4.9953
2.808 | -33.9866 | -5.0059
2.8141 | -33.9869 | -5.0169
2.8202 | -33.9871 | -5.0276
2.8263 | -33.9874 | -5.0384
2.8324 | -33.9878 | -5.0496
2.8385 | -33.9881 | -5.0603
2.8446 | -33.9885 | -5.0709
2.8507 | -33.9888 | -5.0816
2.8568 | -33.9892 | -5.0919
2.8629 | -33.9894 | -5.1023
2.869 | -33.9897 | -5.1133
2.8751 | -33.99 | -5.1239
2.8812 | -33.9904 | -5.135
2.8873 | -33.9907 | -5.1458
2.8934 | -33.991 | -5.1566
2.8995 | -33.9913 | -5.1676
2.9056 | -33.9915 | -5.1786
2.9117 | -33.9919 | -5.1895
2.9178 | -33.9922 | -5.2001
2.9239 | -33.9926 | -5.2115
2.93 | -33.993 | -5.2228
2.9361 | -33.9933 | -5.2335
2.9423 | -33.9936 | -5.244
2.9484 | -33.994 | -5.255
2.9545 | -33.9944 | -5.2658
2.9606 | -33.9948 | -5.2771
2.9667 | -33.9951 | -5.2881
2.9728 | -33.9954 | -5.2993
2.9789 | -33.9958 | -5.3102
2.985 | -33.9963 | -5.3214
2.9911 | -33.9966 | -5.3326
2.9972 | -33.9971 | -5.3435
3.0033 | -33.9974 | -5.3549
3.0094 | -33.9978 | -5.3659
3.0155 | -33.9982 | -5.3767
3.0216 | -33.9986 | -5.3878
3.0277 | -33.9989 | -5.3994
3.0338 | -33.9993 | -5.4107
3.0399 | -33.9997 | -5.4218
3.046 | -34 | -5.4333
3.0521 | -34.0003 | -5.4443
3.0582 | -34.0006 | -5.4563
3.0643 | -34.0009 | -5.4676
3.0704 | -34.0012 | -5.4794
3.0765 | -34.0015 | -5.4903
3.0827 | -34.0018 | -5.5015
3.0888 | -34.0021 | -5.5126
3.0949 | -34.0025 | -5.5239
3.101 | -34.0029 | -5.5352
3.1071 | -34.0032 | -5.5469
3.1132 | -34.0035 | -5.5586
3.1193 | -34.0039 | -5.5699
3.1254 | -34.0043 | -5.5812
3.1315 | -34.0046 | -5.593
3.1376 | -34.0049 | -5.6043
3.1437 | -34.0052 | -5.6158
3.1498 | -34.0056 | -5.6277
3.1559 | -34.0059 | -5.639
3.162 | -34.0061 | -5.6511
3.1681 | -34.0064 | -5.663
3.1742 | -34.0067 | -5.6745
3.1803 | -34.007 | -5.686
3.1864 | -34.0074 | -5.6972
3.1925 | -34.0077 | -5.7087
3.1986 | -34.0079 | -5.7206
3.2047 | -34.0082 | -5.7326
3.2108 | -34.0084 | -5.7436
3.2169 | -34.0087 | -5.7552
3.223 | -34.009 | -5.7669
3.2292 | -34.0092 | -5.7786
3.2353 | -34.0094 | -5.7893
3.2414 | -34.0096 | -5.8007
3.2475 | -34.0098 | -5.812
3.2536 | -34.0101 | -5.8234
3.2597 | -34.0104 | -5.8342
3.2658 | -34.0106 | -5.8451
3.2719 | -34.0108 | -5.856
3.278 | -34.011 | -5.8675
3.2841 | -34.0113 | -5.8785
3.2902 | -34.0115 | -5.8899
3.2963 | -34.0118 | -5.9013
3.3024 | -34.0118 | -5.9123
3.3085 | -34.0121 | -5.9238
3.3146 | -34.0123 | -5.9349
3.3207 | -34.0125 | -5.9462
3.3268 | -34.0127 | -5.9576
3.3329 | -34.0129 | -5.9683
3.339 | -34.0131 | -5.979
3.3451 | -34.0133 | -5.9901
3.3512 | -34.0134 | -6.0007
3.3573 | -34.0136 | -6.0119
3.3634 | -34.0138 | -6.0229
3.3696 | -34.014 | -6.0342
3.3757 | -34.0141 | -6.0449
3.3818 | -34.0144 | -6.0557
3.3879 | -34.0145 | -6.0667
3.394 | -34.0146 | -6.0779
3.4001 | -34.0148 | -6.0887
3.4062 | -34.0149 | -6.0995
3.4123 | -34.0151 | -6.11
3.4184 | -34.0153 | -6.1212
3.4245 | -34.0155 | -6.1319
3.4306 | -34.0155 | -6.1425
3.4367 | -34.0157 | -6.153
3.4428 | -34.0159 | -6.1632
3.4489 | -34.016 | -6.1741
3.455 | -34.0161 | -6.1849
3.4611 | -34.0163 | -6.196
3.4672 | -34.0164 | -6.2069
3.4733 | -34.0166 | -6.2179
3.4794 | -34.0167 | -6.2286
3.4855 | -34.0169 | -6.2389
3.4916 | -34.017 | -6.2492
3.4977 | -34.0172 | -6.2604
3.5038 | -34.0173 | -6.2712
3.5099 | -34.0174 | -6.2814
3.5161 | -34.0176 | -6.2916
3.5222 | -34.0177 | -6.3019
3.5283 | -34.0178 | -6.3126
3.5344 | -34.0179 | -6.3232
3.5405 | -34.018 | -6.3337
3.5466 | -34.0182 | -6.3445
3.5527 | -34.0183 | -6.3552
3.5588 | -34.0184 | -6.3657
3.5649 | -34.0184 | -6.3762
3.571 | -34.0185 | -6.3868
3.5771 | -34.0187 | -6.3978
3.5832 | -34.0188 | -6.4076
3.5893 | -34.019 | -6.4181
3.5954 | -34.0192 | -6.4288
3.6015 | -34.0193 | -6.4396
3.6076 | -34.0195 | -6.45
3.6137 | -34.0196 | -6.46
3.6198 | -34.0199 | -6.4701
3.6259 | -34.02 | -6.4809
3.632 | -34.0202 | -6.4914
3.6381 | -34.0203 | -6.5021
3.6442 | -34.0205 | -6.5125
3.6503 | -34.0207 | -6.5226
3.6565 | -34.0209 | -6.5332
3.6626 | -34.021 | -6.5437
3.6687 | -34.0212 | -6.5543
3.6748 | -34.0214 | -6.5647
3.6809 | -34.0217 | -6.5752
3.687 | -34.0218 | -6.586
3.6931 | -34.0221 | -6.5961
3.6992 | -34.0223 | -6.6067
3.7053 | -34.0225 | -6.617
3.7114 | -34.0226 | -6.6276
3.7175 | -34.0228 | -6.6382
3.7236 | -34.0231 | -6.6481
3.7297 | -34.0233 | -6.6588
3.7358 | -34.0235 | -6.6696
3.7419 | -34.0236 | -6.6803
3.748 | -34.0239 | -6.6908
3.7541 | -34.024 | -6.7013
3.7602 | -34.0243 | -6.7116
3.7663 | -34.0245 | -6.722
3.7724 | -34.0248 | -6.7325
3.7785 | -34.025 | -6.7428
3.7846 | -34.0252 | -6.7531
3.7907 | -34.0255 | -6.764
3.7969 | -34.0258 | -6.7746
3.803 | -34.026 | -6.7847
3.8091 | -34.0262 | -6.7951
3.8152 | -34.0265 | -6.8055
3.8213 | -34.0267 | -6.8162
3.8274 | -34.0269 | -6.8264
3.8335 | -34.0273 | -6.837
3.8396 | -34.0276 | -6.8476
3.8457 | -34.0279 | -6.8582
3.8518 | -34.0282 | -6.8686
3.8579 | -34.0285 | -6.8791
3.864 | -34.0288 | -6.8898
3.8701 | -34.0291 | -6.9006
3.8762 | -34.0293 | -6.911
3.8823 | -34.0296 | -6.9217
3.8884 | -34.0299 | -6.9322
3.8945 | -34.0302 | -6.943
3.9006 | -34.0305 | -6.9539
3.9067 | -34.0308 | -6.9649
3.9128 | -34.0311 | -6.9755
3.9189 | -34.0314 | -6.9858
3.925 | -34.0316 | -6.9965
3.9311 | -34.0319 | -7.0066
3.9372 | -34.0322 | -7.0174
3.9434 | -34.0325 | -7.028
3.9495 | -34.0328 | -7.0387
3.9556 | -34.0332 | -7.0491
3.9617 | -34.0334 | -7.0593
3.9678 | -34.0338 | -7.0696
3.9739 | -34.034 | -7.0801
3.98 | -34.0343 | -7.0908
3.9861 | -34.0346 | -7.1012
3.9922 | -34.0348 | -7.1118
3.9983 | -34.0351 | -7.122
4.0044 | -34.0354 | -7.1321
4.0105 | -34.0357 | -7.1427
4.0166 | -34.036 | -7.1533
4.0227 | -34.0363 | -7.1638
4.0288 | -34.0366 | -7.1748
4.0349 | -34.0368 | -7.1851
4.041 | -34.037 | -7.1956
4.0471 | -34.0373 | -7.2062
4.0532 | -34.0376 | -7.2164
4.0593 | -34.0379 | -7.2271
4.0654 | -34.0382 | -7.2374
4.0715 | -34.0384 | -7.2484
4.0776 | -34.0388 | -7.2588
4.0838 | -34.0391 | -7.2697
4.0899 | -34.0393 | -7.28
4.096 | -34.0396 | -7.2904
4.1021 | -34.0398 | -7.3007
4.1082 | -34.0401 | -7.311
4.1143 | -34.0404 | -7.3212
4.1204 | -34.0407 | -7.3314
4.1265 | -34.0409 | -7.3416
4.1326 | -34.0412 | -7.3517
4.1387 | -34.0415 | -7.362
4.1448 | -34.0418 | -7.3725
4.1509 | -34.042 | -7.383
4.157 | -34.0422 | -7.3938
4.1631 | -34.0425 | -7.4045
4.1692 | -34.0428 | -7.4147
4.1753 | -34.0431 | -7.4251
4.1814 | -34.0433 | -7.4352
4.1875 | -34.0436 | -7.4459
4.1936 | -34.0438 | -7.4564
4.1997 | -34.044 | -7.4665
4.2058 | -34.0443 | -7.4772
4.2119 | -34.0445 | -7.4874
4.218 | -34.0447 | -7.4975
4.2241 | -34.045 | -7.5077
4.2303 | -34.0453 | -7.5179
4.2364 | -34.0455 | -7.5283
4.2425 | -34.0457 | -7.5386
4.2486 | -34.0459 | -7.5492
4.2547 | -34.046 | -7.5593
4.2608 | -34.0462 | -7.5698
4.2669 | -34.0464 | -7.5797
4.273 | -34.0467 | -7.5902
4.2791 | -34.0469 | -7.6006
4.2852 | -34.0471 | -7.611
4.2913 | -34.0473 | -7.6211
4.2974 | -34.0475 | -7.6316
4.3035 | -34.0477 | -7.6423
4.3096 | -34.0479 | -7.6528
4.3157 | -34.0482 | -7.6629
4.3218 | -34.0484 | -7.6731
4.3279 | -34.0486 | -7.6834
4.334 | -34.0488 | -7.6936
4.3401 | -34.049 | -7.7035
4.3462 | -34.0492 | -7.7139
4.3523 | -34.0495 | -7.7237
4.3584 | -34.0496 | -7.7336
4.3645 | -34.0498 | -7.7437
4.3707 | -34.05 | -7.7537
4.3768 | -34.0502 | -7.7639
4.3829 | -34.0505 | -7.7745
4.389 | -34.0507 | -7.7847
4.3951 | -34.0508 | -7.7947
4.4012 | -34.051 | -7.8052
4.4073 | -34.0512 | -7.8153
4.4134 | -34.0514 | -7.8254
4.4195 | -34.0517 | -7.8355
4.4256 | -34.0519 | -7.8455
4.4317 | -34.0521 | -7.8557
4.4378 | -34.0523 | -7.8655
4.4439 | -34.0525 | -7.8754
4.45 | -34.0526 | -7.8855
4.4561 | -34.0528 | -7.8957
4.4622 | -34.0529 | -7.9059
4.4683 | -34.0532 | -7.916
4.4744 | -34.0533 | -7.9265
4.4805 | -34.0535 | -7.9362
4.4866 | -34.0537 | -7.9463
4.4927 | -34.054 | -7.9567
4.4988 | -34.0541 | -7.9668
4.5049 | -34.0544 | -7.9765
4.511 | -34.0546 | -7.9863
4.5172 | -34.0548 | -7.9964
4.5233 | -34.0549 | -8.0069
4.5294 | -34.0552 | -8.0168
4.5355 | -34.0553 | -8.0269
4.5416 | -34.0555 | -8.037
4.5477 | -34.0557 | -8.0471
4.5538 | -34.0559 | -8.0568
4.5599 | -34.0561 | -8.0668
4.566 | -34.0563 | -8.0769
4.5721 | -34.0565 | -8.087
4.5782 | -34.0568 | -8.0973
4.5843 | -34.057 | -8.1075
4.5904 | -34.0572 | -8.1176
4.5965 | -34.0574 | -8.1276
4.6026 | -34.0576 | -8.1381
4.6087 | -34.0578 | -8.1481
4.6148 | -34.058 | -8.1585
4.6209 | -34.0582 | -8.1686
4.627 | -34.0584 | -8.1789
4.6331 | -34.0586 | -8.1888
4.6392 | -34.0588 | -8.1991
4.6453 | -34.059 | -8.2091
4.6514 | -34.0592 | -8.2193
4.6576 | -34.0594 | -8.2297
4.6637 | -34.0595 | -8.2395
4.6698 | -34.0597 | -8.2494
4.6759 | -34.0599 | -8.2599
4.682 | -34.0601 | -8.2699
4.6881 | -34.0602 | -8.28
4.6942 | -34.0605 | -8.2901
4.7003 | -34.0606 | -8.3003
4.7064 | -34.0609 | -8.3101
4.7125 | -34.0611 | -8.3205
4.7186 | -34.0613 | -8.3304
4.7247 | -34.0615 | -8.3406
4.7308 | -34.0617 | -8.3506
4.7369 | -34.0619 | -8.3608
4.743 | -34.0621 | -8.3713
4.7491 | -34.0623 | -8.3818
4.7552 | -34.0625 | -8.392
4.7613 | -34.0627 | -8.4023
4.7674 | -34.0628 | -8.4122
4.7735 | -34.063 | -8.4228
4.7796 | -34.0633 | -8.4333
4.7857 | -34.0635 | -8.4435
4.7918 | -34.0637 | -8.4535
4.7979 | -34.064 | -8.4636
4.8041 | -34.0642 | -8.4738
4.8102 | -34.0644 | -8.4836
4.8163 | -34.0646 | -8.4938
4.8224 | -34.0648 | -8.504
4.8285 | -34.065 | -8.5138
4.8346 | -34.0653 | -8.5238
4.8407 | -34.0654 | -8.534
4.8468 | -34.0657 | -8.5445
4.8529 | -34.0659 | -8.5549
4.859 | -34.0661 | -8.5654
4.8651 | -34.0663 | -8.5753
4.8712 | -34.0665 | -8.5853
4.8773 | -34.0667 | -8.5956
4.8834 | -34.0669 | -8.6058
4.8895 | -34.0672 | -8.6157
4.8956 | -34.0673 | -8.6259
4.9017 | -34.0675 | -8.6359
4.9078 | -34.0677 | -8.6464
4.9139 | -34.0679 | -8.6561
4.92 | -34.0682 | -8.666
4.9261 | -34.0685 | -8.6759
4.9322 | -34.0687 | -8.6861
4.9383 | -34.0689 | -8.6962
4.9445 | -34.0691 | -8.7064
4.9506 | -34.0693 | -8.7164
4.9567 | -34.0696 | -8.7263
4.9628 | -34.0698 | -8.7364
4.9689 | -34.07 | -8.747
4.975 | -34.0702 | -8.7574
4.9811 | -34.0704 | -8.7673
4.9872 | -34.0707 | -8.7774
4.9933 | -34.071 | -8.7878
4.9994 | -34.0712 | -8.7984
5.0055 | -34.0715 | -8.8086
5.0116 | -34.0717 | -8.8192
5.0177 | -34.072 | -8.8297
5.0238 | -34.0723 | -8.8402
5.0299 | -34.0726 | -8.8508
5.036 | -34.0728 | -8.861
5.0421 | -34.073 | -8.8711
5.0482 | -34.0733 | -8.8814
5.0543 | -34.0735 | -8.8913
5.0604 | -34.0738 | -8.9015
5.0665 | -34.0741 | -8.9115
5.0726 | -34.0743 | -8.922
5.0787 | -34.0746 | -8.9325
5.0848 | -34.0748 | -8.9428
5.091 | -34.0751 | -8.9528
5.0971 | -34.0754 | -8.9627
5.1032 | -34.0756 | -8.9733
5.1093 | -34.0759 | -8.9835
5.1154 | -34.0761 | -8.9935
5.1215 | -34.0763 | -9.0035
5.1276 | -34.0766 | -9.0135
5.1337 | -34.0769 | -9.0235
5.1398 | -34.0771 | -9.0334
5.1459 | -34.0774 | -9.0439
5.152 | -34.0776 | -9.0543
5.1581 | -34.0779 | -9.0647
5.1642 | -34.0782 | -9.075
5.1703 | -34.0784 | -9.0853
5.1764 | -34.0787 | -9.0954
5.1825 | -34.079 | -9.1053
5.1886 | -34.0792 | -9.1153
5.1947 | -34.0795 | -9.1255
5.2008 | -34.0798 | -9.1359
5.2069 | -34.08 | -9.1464
5.213 | -34.0803 | -9.1564
5.2191 | -34.0806 | -9.1666
5.2252 | -34.0808 | -9.1769
5.2314 | -34.0811 | -9.1873
5.2375 | -34.0813 | -9.1973
5.2436 | -34.0816 | -9.2077
5.2497 | -34.0819 | -9.2176
5.2558 | -34.0821 | -9.2282
5.2619 | -34.0824 | -9.2388
5.268 | -34.0827 | -9.2491
5.2741 | -34.083 | -9.2595
5.2802 | -34.0833 | -9.2701
5.2863 | -34.0836 | -9.2803
5.2924 | -34.0839 | -9.2908
5.2985 | -34.0842 | -9.3011
5.3046 | -34.0844 | -9.3114
5.3107 | -34.0847 | -9.322
5.3168 | -34.085 | -9.3323
5.3229 | -34.0853 | -9.3426
5.329 | -34.0855 | -9.3529
5.3351 | -34.0858 | -9.3629
5.3412 | -34.0861 | -9.3732
5.3473 | -34.0864 | -9.3835
5.3534 | -34.0867 | -9.394
5.3595 | -34.0869 | -9.4043
5.3656 | -34.0873 | -9.4146
5.3717 | -34.0875 | -9.4253
5.3779 | -34.0878 | -9.4359
5.384 | -34.0881 | -9.4462
5.3901 | -34.0884 | -9.4568
5.3962 | -34.0887 | -9.4669
5.4023 | -34.089 | -9.4775
5.4084 | -34.0893 | -9.4877
5.4145 | -34.0896 | -9.4978
5.4206 | -34.0899 | -9.5082
5.4267 | -34.0902 | -9.5188
5.4328 | -34.0904 | -9.5292
5.4389 | -34.0907 | -9.5393
5.445 | -34.091 | -9.55
5.4511 | -34.0913 | -9.5601
5.4572 | -34.0916 | -9.5706
5.4633 | -34.0919 | -9.5809
5.4694 | -34.0921 | -9.5913
5.4755 | -34.0924 | -9.6014
5.4816 | -34.0927 | -9.6117
5.4877 | -34.093 | -9.6219
5.4938 | -34.0933 | -9.632
5.4999 | -34.0936 | -9.6423
5.506 | -34.094 | -9.6526
5.5121 | -34.0942 | -9.6628
5.5183 | -34.0945 | -9.673
5.5244 | -34.0948 | -9.6836
5.5305 | -34.0951 | -9.6938
5.5366 | -34.0954 | -9.7039
5.5427 | -34.0956 | -9.7144
5.5488 | -34.0959 | -9.7248
5.5549 | -34.0962 | -9.7352
5.561 | -34.0965 | -9.7457
5.5671 | -34.0968 | -9.7559
5.5732 | -34.0971 | -9.7665
5.5793 | -34.0974 | -9.7772
5.5854 | -34.0977 | -9.7875
5.5915 | -34.098 | -9.7982
5.5976 | -34.0983 | -9.8086
5.6037 | -34.0986 | -9.819
5.6098 | -34.0988 | -9.8293
5.6159 | -34.0992 | -9.8395
5.622 | -34.0994 | -9.85
5.6281 | -34.0997 | -9.8604
5.6342 | -34.1 | -9.8711
5.6403 | -34.1003 | -9.8815
5.6464 | -34.1006 | -9.892
5.6525 | -34.1009 | -9.9025
5.6586 | -34.1012 | -9.9125
5.6648 | -34.1016 | -9.9229
5.6709 | -34.1019 | -9.9335
5.677 | -34.1022 | -9.9439
5.6831 | -34.1025 | -9.9543
5.6892 | -34.1027 | -9.9649
5.6953 | -34.103 | -9.9753
5.7014 | -34.1034 | -9.9858
5.7075 | -34.1036 | -9.9961
5.7136 | -34.1039 | -10.0065
5.7197 | -34.1042 | -10.0168
5.7258 | -34.1044 | -10.027
5.7319 | -34.1048 | -10.0375
5.738 | -34.1051 | -10.0478
5.7441 | -34.1054 | -10.0582
5.7502 | -34.1057 | -10.0682
5.7563 | -34.106 | -10.0786
5.7624 | -34.1063 | -10.0892
5.7685 | -34.1067 | -10.0996
5.7746 | -34.1069 | -10.1102
5.7807 | -34.1072 | -10.1207
5.7868 | -34.1074 | -10.1313
5.7929 | -34.1077 | -10.1415
5.799 | -34.108 | -10.1518
5.8052 | -34.1083 | -10.162
5.8113 | -34.1086 | -10.1723
5.8174 | -34.1089 | -10.1825
5.8235 | -34.1092 | -10.1928
5.8296 | -34.1094 | -10.2034
5.8357 | -34.1097 | -10.2137
5.8418 | -34.11 | -10.2243
5.8479 | -34.1103 | -10.2347
5.854 | -34.1106 | -10.2452
5.8601 | -34.1109 | -10.2554
5.8662 | -34.1112 | -10.2659
5.8723 | -34.1114 | -10.2761
5.8784 | -34.1117 | -10.2864
5.8845 | -34.112 | -10.2967
5.8906 | -34.1123 | -10.3071
5.8967 | -34.1126 | -10.3172
5.9028 | -34.1129 | -10.3274
5.9089 | -34.1132 | -10.3376
5.915 | -34.1135 | -10.3478
5.9211 | -34.1138 | -10.3581
5.9272 | -34.1141 | -10.3682
5.9333 | -34.1144 | -10.3784
5.9394 | -34.1147 | -10.3887
5.9455 | -34.115 | -10.3988
5.9517 | -34.1153 | -10.4089
5.9578 | -34.1157 | -10.4194
5.9639 | -34.1159 | -10.4293
5.97 | -34.1163 | -10.4397
5.9761 | -34.1166 | -10.45
5.9822 | -34.1169 | -10.4603
5.9883 | -34.1172 | -10.4706
5.9944 | -34.1175 | -10.481
6.0005 | -34.1178 | -10.4915
6.0066 | -34.1181 | -10.5016
6.0127 | -34.1184 | -10.5116
6.0188 | -34.1187 | -10.5222
6.0249 | -34.119 | -10.5326
6.031 | -34.1193 | -10.5427
6.0371 | -34.1196 | -10.553
6.0432 | -34.1199 | -10.5631
6.0493 | -34.1202 | -10.5732
6.0554 | -34.1205 | -10.5833
6.0615 | -34.1208 | -10.5936
6.0676 | -34.1211 | -10.6041
6.0737 | -34.1215 | -10.6145
6.0798 | -34.1217 | -10.6247
6.0859 | -34.1221 | -10.6349
6.0921 | -34.1224 | -10.6449
6.0982 | -34.1227 | -10.6552
6.1043 | -34.123 | -10.6653
6.1104 | -34.1233 | -10.6758
6.1165 | -34.1236 | -10.6856
6.1226 | -34.124 | -10.6957
6.1287 | -34.1242 | -10.7059
6.1348 | -34.1246 | -10.7161
6.1409 | -34.1249 | -10.7263
6.147 | -34.1252 | -10.7364
6.1531 | -34.1254 | -10.7466
6.1592 | -34.1257 | -10.7567
6.1653 | -34.126 | -10.7667
6.1714 | -34.1263 | -10.7769
6.1775 | -34.1267 | -10.7873
6.1836 | -34.1269 | -10.7975
6.1897 | -34.1273 | -10.8077
6.1958 | -34.1276 | -10.8181
6.2019 | -34.1279 | -10.8284
6.208 | -34.1282 | -10.8386
6.2141 | -34.1285 | -10.8487
6.2202 | -34.1289 | -10.859
6.2263 | -34.1292 | -10.8691
6.2325 | -34.1294 | -10.879
6.2386 | -34.1298 | -10.8889
6.2447 | -34.1301 | -10.8993
6.2508 | -34.1304 | -10.9094
6.2569 | -34.1307 | -10.9195
6.263 | -34.131 | -10.9297
6.2691 | -34.1313 | -10.9398
6.2752 | -34.1316 | -10.95
6.2813 | -34.1319 | -10.9601
6.2874 | -34.1322 | -10.9701
6.2935 | -34.1325 | -10.9802
6.2996 | -34.1328 | -10.9904
6.3057 | -34.1331 | -11.0006
6.3118 | -34.1333 | -11.0108
6.3179 | -34.1336 | -11.0209
6.324 | -34.134 | -11.031
6.3301 | -34.1343 | -11.0411
6.3362 | -34.1346 | -11.0511
6.3423 | -34.1349 | -11.0614
6.3484 | -34.1352 | -11.0716
6.3545 | -34.1355 | -11.0814
6.3606 | -34.1358 | -11.0912
6.3667 | -34.136 | -11.1012
6.3728 | -34.1364 | -11.1113
6.379 | -34.1367 | -11.1214
6.3851 | -34.137 | -11.1313
6.3912 | -34.1373 | -11.1413
6.3973 | -34.1375 | -11.1515
6.4034 | -34.1378 | -11.1613
6.4095 | -34.1381 | -11.1714
6.4156 | -34.1384 | -11.1815
6.4217 | -34.1387 | -11.1915
6.4278 | -34.139 | -11.2017
6.4339 | -34.1394 | -11.2117
6.44 | -34.1397 | -11.2217
6.4461 | -34.14 | -11.232
6.4522 | -34.1403 | -11.2418
6.4583 | -34.1405 | -11.252
6.4644 | -34.1408 | -11.2622
6.4705 | -34.1411 | -11.2719
6.4766 | -34.1414 | -11.282
6.4827 | -34.1417 | -11.2921
6.4888 | -34.142 | -11.3019
6.4949 | -34.1423 | -11.3117
6.501 | -34.1426 | -11.3218
6.5071 | -34.1429 | -11.3316
6.5132 | -34.1432 | -11.3417
6.5194 | -34.1435 | -11.3518
6.5255 | -34.1438 | -11.3615
6.5316 | -34.1441 | -11.3717
6.5377 | -34.1443 | -11.3817
6.5438 | -34.1446 | -11.3915
6.5499 | -34.1449 | -11.4014
6.556 | -34.1452 | -11.4114
6.5621 | -34.1455 | -11.4216
6.5682 | -34.1458 | -11.4317
6.5743 | -34.1462 | -11.4416
6.5804 | -34.1464 | -11.4517
6.5865 | -34.1467 | -11.4617
6.5926 | -34.1471 | -11.4716
6.5987 | -34.1473 | -11.4813
6.6048 | -34.1476 | -11.4913
6.6109 | -34.1479 | -11.5013
6.617 | -34.1482 | -11.5114
6.6231 | -34.1485 | -11.5213
6.6292 | -34.1488 | -11.5313
6.6353 | -34.1491 | -11.5416
6.6414 | -34.1494 | -11.5514
6.6475 | -34.1497 | -11.5614
6.6536 | -34.15 | -11.5713
6.6597 | -34.1503 | -11.5809
6.6659 | -34.1506 | -11.5909
6.672 | -34.1509 | -11.6007
6.6781 | -34.1512 | -11.6107
6.6842 | -34.1515 | -11.6205
6.6903 | -34.1519 | -11.6304
6.6964 | -34.1522 | -11.6404
6.7025 | -34.1524 | -11.6501
6.7086 | -34.1527 | -11.66
6.7147 | -34.153 | -11.6701
6.7208 | -34.1533 | -11.6801
6.7269 | -34.1536 | -11.69
6.733 | -34.1539 | -11.6997
6.7391 | -34.1542 | -11.7096
6.7452 | -34.1545 | -11.7193
6.7513 | -34.1548 | -11.729
6.7574 | -34.1551 | -11.739
6.7635 | -34.1554 | -11.7487
6.7696 | -34.1557 | -11.7585
6.7757 | -34.156 | -11.7684
6.7818 | -34.1563 | -11.7783
6.7879 | -34.1567 | -11.788
6.794 | -34.157 | -11.7977
6.8001 | -34.1572 | -11.8075
6.8063 | -34.1575 | -11.8174
6.8124 | -34.1578 | -11.8274
6.8185 | -34.1581 | -11.8374
6.8246 | -34.1584 | -11.8471
6.8307 | -34.1587 | -11.8569
6.8368 | -34.159 | -11.867
6.8429 | -34.1593 | -11.8768
6.849 | -34.1596 | -11.8868
6.8551 | -34.16 | -11.8965
6.8612 | -34.1602 | -11.9063
6.8673 | -34.1605 | -11.9161
6.8734 | -34.1608 | -11.9261
6.8795 | -34.1611 | -11.936
6.8856 | -34.1614 | -11.9458
6.8917 | -34.1617 | -11.9556
6.8978 | -34.162 | -11.9651
6.9039 | -34.1623 | -11.9747
6.91 | -34.1626 | -11.9846
6.9161 | -34.1629 | -11.9942
6.9222 | -34.1632 | -12.004
6.9283 | -34.1636 | -12.0136
6.9344 | -34.1639 | -12.0234
6.9405 | -34.1642 | -12.0332
6.9466 | -34.1646 | -12.0432
6.9528 | -34.1649 | -12.0532
6.9589 | -34.1652 | -12.0632
6.965 | -34.1655 | -12.0731
6.9711 | -34.1658 | -12.083
6.9772 | -34.1662 | -12.0928
6.9833 | -34.1665 | -12.1026
6.9894 | -34.1668 | -12.1127
6.9955 | -34.1671 | -12.1222
7.0016 | -34.1674 | -12.132
7.0077 | -34.1677 | -12.1418
7.0138 | -34.168 | -12.1515
7.0199 | -34.1683 | -12.1612
7.026 | -34.1686 | -12.171
7.0321 | -34.169 | -12.1806
7.0382 | -34.1693 | -12.19
7.0443 | -34.1696 | -12.2
7.0504 | -34.1699 | -12.2099
7.0565 | -34.1702 | -12.2197
7.0626 | -34.1706 | -12.2292
7.0687 | -34.1709 | -12.2385
7.0748 | -34.1712 | -12.248
7.0809 | -34.1715 | -12.2578
7.087 | -34.1718 | -12.2676
7.0932 | -34.1721 | -12.2774
7.0993 | -34.1724 | -12.2871
7.1054 | -34.1728 | -12.2967
7.1115 | -34.1731 | -12.3062
7.1176 | -34.1734 | -12.3159
7.1237 | -34.1737 | -12.3256
7.1298 | -34.174 | -12.3352
7.1359 | -34.1743 | -12.345
7.142 | -34.1747 | -12.3544
7.1481 | -34.175 | -12.364
7.1542 | -34.1753 | -12.3738
7.1603 | -34.1756 | -12.3835
7.1664 | -34.1759 | -12.3931
7.1725 | -34.1763 | -12.403
7.1786 | -34.1766 | -12.4126
7.1847 | -34.1769 | -12.4224
7.1908 | -34.1772 | -12.4323
7.1969 | -34.1775 | -12.4421
7.203 | -34.1778 | -12.4517
7.2091 | -34.1782 | -12.4613
7.2152 | -34.1785 | -12.471
7.2213 | -34.1788 | -12.4806
7.2274 | -34.1792 | -12.4902
7.2335 | -34.1795 | -12.4998
7.2397 | -34.1798 | -12.5095
7.2458 | -34.1802 | -12.5189
7.2519 | -34.1805 | -12.5287
7.258 | -34.1808 | -12.5385
7.2641 | -34.1811 | -12.5482
7.2702 | -34.1814 | -12.5581
7.2763 | -34.1818 | -12.5678
7.2824 | -34.1821 | -12.5776
7.2885 | -34.1824 | -12.5874
7.2946 | -34.1827 | -12.597
7.3007 | -34.183 | -12.6068
7.3068 | -34.1833 | -12.6168
7.3129 | -34.1837 | -12.6266
7.319 | -34.184 | -12.6364
7.3251 | -34.1843 | -12.6459
7.3312 | -34.1847 | -12.6556
7.3373 | -34.185 | -12.665
7.3434 | -34.1854 | -12.6746
7.3495 | -34.1857 | -12.6842
7.3556 | -34.186 | -12.694
7.3617 | -34.1864 | -12.7039
7.3678 | -34.1867 | -12.7135
7.3739 | -34.187 | -12.7234
7.3801 | -34.1873 | -12.7333
7.3862 | -34.1877 | -12.743
7.3923 | -34.188 | -12.7528
7.3984 | -34.1883 | -12.7623
7.4045 | -34.1887 | -12.7719
7.4106 | -34.189 | -12.7816
7.4167 | -34.1893 | -12.7913
7.4228 | -34.1897 | -12.8012
7.4289 | -34.19 | -12.8108
7.435 | -34.1904 | -12.8203
7.4411 | -34.1907 | -12.8301
7.4472 | -34.191 | -12.8399
7.4533 | -34.1914 | -12.8494
7.4594 | -34.1917 | -12.8593
7.4655 | -34.192 | -12.8688
7.4716 | -34.1924 | -12.8787
7.4777 | -34.1927 | -12.8884
7.4838 | -34.193 | -12.8979
7.4899 | -34.1934 | -12.9076
7.496 | -34.1937 | -12.9171
7.5021 | -34.194 | -12.9268
7.5082 | -34.1944 | -12.9365
7.5143 | -34.1947 | -12.9462
7.5204 | -34.195 | -12.956
7.5266 | -34.1954 | -12.9656
7.5327 | -34.1957 | -12.9752
7.5388 | -34.196 | -12.9853
7.5449 | -34.1964 | -12.9951
7.551 | -34.1967 | -13.0047
7.5571 | -34.1971 | -13.0145
7.5632 | -34.1974 | -13.0243
7.5693 | -34.1978 | -13.0339
7.5754 | -34.1981 | -13.0435
7.5815 | -34.1984 | -13.0532
7.5876 | -34.1988 | -13.0631
7.5937 | -34.1991 | -13.073
7.5998 | -34.1995 | -13.0826
7.6059 | -34.1999 | -13.0925
7.612 | -34.2002 | -13.1022
7.6181 | -34.2005 | -13.1118
7.6242 | -34.2009 | -13.1218
7.6303 | -34.2012 | -13.1318
7.6364 | -34.2015 | -13.1416
7.6425 | -34.2019 | -13.1513
7.6486 | -34.2022 | -13.1611
7.6547 | -34.2025 | -13.1707
7.6608 | -34.2029 | -13.1804
7.667 | -34.2032 | -13.1903
7.6731 | -34.2036 | -13.2001
7.6792 | -34.2039 | -13.2101
7.6853 | -34.2043 | -13.2199
7.6914 | -34.2046 | -13.2297
7.6975 | -34.2049 | -13.2396
7.7036 | -34.2053 | -13.2494
7.7097 | -34.2057 | -13.2593
7.7158 | -34.206 | -13.2692
7.7219 | -34.2064 | -13.2788
7.728 | -34.2067 | -13.2887
7.7341 | -34.207 | -13.2986
7.7402 | -34.2074 | -13.3085
7.7463 | -34.2077 | -13.3183
7.7524 | -34.2081 | -13.3284
7.7585 | -34.2084 | -13.3382
7.7646 | -34.2088 | -13.3479
7.7707 | -34.2091 | -13.3579
7.7768 | -34.2094 | -13.3679
7.7829 | -34.2098 | -13.3779
7.789 | -34.2101 | -13.388
7.7951 | -34.2104 | -13.3978
7.8012 | -34.2108 | -13.4076
7.8073 | -34.2112 | -13.4177
7.8135 | -34.2115 | -13.4275
7.8196 | -34.2119 | -13.4375
7.8257 | -34.2122 | -13.4476
7.8318 | -34.2125 | -13.4575
7.8379 | -34.2129 | -13.4672
7.844 | -34.2132 | -13.477
7.8501 | -34.2136 | -13.4869
7.8562 | -34.2139 | -13.497
7.8623 | -34.2143 | -13.5067
7.8684 | -34.2147 | -13.5167
7.8745 | -34.215 | -13.5266
7.8806 | -34.2153 | -13.5367
7.8867 | -34.2157 | -13.5463
7.8928 | -34.2161 | -13.5559
7.8989 | -34.2164 | -13.5658
7.905 | -34.2168 | -13.5757
7.9111 | -34.2171 | -13.5857
7.9172 | -34.2174 | -13.5955
7.9233 | -34.2178 | -13.6055
7.9294 | -34.2181 | -13.6154
7.9355 | -34.2185 | -13.625
7.9416 | -34.2188 | -13.6349
7.9477 | -34.2191 | -13.6447
7.9539 | -34.2195 | -13.6545
7.96 | -34.2198 | -13.6645
7.9661 | -34.2202 | -13.6744
7.9722 | -34.2206 | -13.6841
7.9783 | -34.2209 | -13.6942
7.9844 | -34.2213 | -13.7041
7.9905 | -34.2216 | -13.7142
7.9966 | -34.222 | -13.7242
8.0027 | -34.2223 | -13.7341
8.0088 | -34.2227 | -13.7439
8.0149 | -34.223 | -13.7539
8.021 | -34.2233 | -13.7639
8.0271 | -34.2237 | -13.774
8.0332 | -34.224 | -13.7839
8.0393 | -34.2244 | -13.7939
8.0454 | -34.2248 | -13.8041
8.0515 | -34.2252 | -13.814
8.0576 | -34.2255 | -13.8241
8.0637 | -34.2258 | -13.8339
8.0698 | -34.2262 | -13.8437
8.0759 | -34.2265 | -13.8535
8.082 | -34.2268 | -13.8633
8.0881 | -34.2272 | -13.8733
8.0942 | -34.2276 | -13.8832
8.1004 | -34.228 | -13.8932
8.1065 | -34.2283 | -13.903
8.1126 | -34.2286 | -13.9129
8.1187 | -34.229 | -13.9229
8.1248 | -34.2293 | -13.9326
8.1309 | -34.2297 | -13.9423
8.137 | -34.23 | -13.9521
8.1431 | -34.2304 | -13.9622
8.1492 | -34.2307 | -13.9721
8.1553 | -34.2311 | -13.9821
8.1614 | -34.2314 | -13.9922
8.1675 | -34.2318 | -14.0021
8.1736 | -34.2321 | -14.0122
8.1797 | -34.2324 | -14.0222
8.1858 | -34.2328 | -14.0322
8.1919 | -34.2331 | -14.0421
8.198 | -34.2335 | -14.0518
8.2041 | -34.2339 | -14.0617
8.2102 | -34.2342 | -14.0716
8.2163 | -34.2346 | -14.0813
8.2224 | -34.2349 | -14.0912
8.2285 | -34.2353 | -14.1009
8.2346 | -34.2356 | -14.1107
8.2408 | -34.236 | -14.1205
8.2469 | -34.2363 | -14.1305
8.253 | -34.2367 | -14.1402
8.2591 | -34.2371 | -14.1502
8.2652 | -34.2375 | -14.1602
8.2713 | -34.2378 | -14.1701
8.2774 | -34.2382 | -14.1801
8.2835 | -34.2385 | -14.1903
8.2896 | -34.2389 | -14.2
8.2957 | -34.2392 | -14.21
8.3018 | -34.2396 | -14.2199
8.3079 | -34.2399 | -14.2296
8.314 | -34.2403 | -14.2395
8.3201 | -34.2407 | -14.2493
8.3262 | -34.241 | -14.2592
8.3323 | -34.2414 | -14.2691
8.3384 | -34.2417 | -14.2791
8.3445 | -34.2421 | -14.2889
8.3506 | -34.2424 | -14.2989
8.3567 | -34.2428 | -14.3088
8.3628 | -34.2431 | -14.3187
8.3689 | -34.2435 | -14.3284
8.375 | -34.2438 | -14.3384
8.3812 | -34.2442 | -14.3482
8.3873 | -34.2446 | -14.3581
8.3934 | -34.2449 | -14.3681
8.3995 | -34.2452 | -14.3778
8.4056 | -34.2456 | -14.3877
8.4117 | -34.2459 | -14.3977
8.4178 | -34.2463 | -14.4076
8.4239 | -34.2466 | -14.4173
8.43 | -34.247 | -14.4272
8.4361 | -34.2473 | -14.4372
8.4422 | -34.2477 | -14.447
8.4483 | -34.248 | -14.4568
8.4544 | -34.2484 | -14.4667
8.4605 | -34.2488 | -14.4764
8.4666 | -34.2491 | -14.4866
8.4727 | -34.2494 | -14.4965
8.4788 | -34.2498 | -14.5063
8.4849 | -34.2501 | -14.5162
8.491 | -34.2505 | -14.5261
8.4971 | -34.2509 | -14.5357
8.5032 | -34.2512 | -14.5454
8.5093 | -34.2516 | -14.5551
8.5154 | -34.2519 | -14.565
8.5215 | -34.2523 | -14.5747
8.5277 | -34.2526 | -14.5846
8.5338 | -34.253 | -14.5945
8.5399 | -34.2533 | -14.6043
8.546 | -34.2537 | -14.6139
8.5521 | -34.2541 | -14.6237
8.5582 | -34.2544 | -14.6335
8.5643 | -34.2547 | -14.6431
8.5704 | -34.2551 | -14.6526
8.5765 | -34.2555 | -14.6626
8.5826 | -34.2558 | -14.6721
8.5887 | -34.2562 | -14.6818
8.5948 | -34.2566 | -14.6917
8.6009 | -34.2569 | -14.7013
8.607 | -34.2573 | -14.7113
8.6131 | -34.2576 | -14.7211
8.6192 | -34.258 | -14.7307
8.6253 | -34.2584 | -14.7404
8.6314 | -34.2587 | -14.7501
8.6375 | -34.259 | -14.7599
8.6436 | -34.2594 | -14.7697
8.6497 | -34.2598 | -14.7793
8.6558 | -34.2601 | -14.789
8.6619 | -34.2604 | -14.7988
8.6681 | -34.2608 | -14.8088
8.6742 | -34.2612 | -14.8185
8.6803 | -34.2615 | -14.8283
8.6864 | -34.2619 | -14.8377
8.6925 | -34.2622 | -14.8474
8.6986 | -34.2626 | -14.8572
8.7047 | -34.2629 | -14.8672
8.7108 | -34.2632 | -14.877
8.7169 | -34.2636 | -14.8867
8.723 | -34.264 | -14.8962
8.7291 | -34.2643 | -14.9058
8.7352 | -34.2647 | -14.9158
8.7413 | -34.265 | -14.9256
8.7474 | -34.2654 | -14.9352
8.7535 | -34.2657 | -14.9449
8.7596 | -34.2661 | -14.9546
8.7657 | -34.2665 | -14.9644
8.7718 | -34.2668 | -14.9742
8.7779 | -34.2672 | -14.9841
8.784 | -34.2675 | -14.9939
8.7901 | -34.2679 | -15.0036
8.7962 | -34.2683 | -15.0131
8.8023 | -34.2686 | -15.0227
8.8084 | -34.269 | -15.0324
8.8146 | -34.2693 | -15.0421
8.8207 | -34.2697 | -15.0517
8.8268 | -34.2701 | -15.0611
8.8329 | -34.2704 | -15.0708
8.839 | -34.2708 | -15.0806
8.8451 | -34.2711 | -15.0902
8.8512 | -34.2715 | -15.0998
8.8573 | -34.2718 | -15.1095
8.8634 | -34.2722 | -15.119
8.8695 | -34.2726 | -15.1289
8.8756 | -34.2729 | -15.1386
8.8817 | -34.2732 | -15.1481
8.8878 | -34.2736 | -15.158
8.8939 | -34.274 | -15.1677
8.9 | -34.2744 | -15.1774
8.9061 | -34.2747 | -15.187
8.9122 | -34.2751 | -15.1967
8.9183 | -34.2754 | -15.2065
8.9244 | -34.2758 | -15.2162
8.9305 | -34.2762 | -15.2257
8.9366 | -34.2766 | -15.2353
8.9427 | -34.2769 | -15.2451
8.9488 | -34.2773 | -15.2547
8.955 | -34.2776 | -15.2643
8.9611 | -34.278 | -15.2741
8.9672 | -34.2784 | -15.2836
8.9733 | -34.2788 | -15.2931
8.9794 | -34.2792 | -15.3026
8.9855 | -34.2795 | -15.3124
8.9916 | -34.2799 | -15.3222
8.9977 | -34.2802 | -15.3318
9.0038 | -34.2807 | -15.3413
9.0099 | -34.281 | -15.351
9.016 | -34.2814 | -15.3607
9.0221 | -34.2818 | -15.3701
9.0282 | -34.2821 | -15.3796
9.0343 | -34.2825 | -15.3892
9.0404 | -34.2829 | -15.3989
9.0465 | -34.2833 | -15.4086
9.0526 | -34.2837 | -15.4182
9.0587 | -34.2841 | -15.4278
9.0648 | -34.2845 | -15.4373
9.0709 | -34.2848 | -15.4469
9.077 | -34.2852 | -15.4566
9.0831 | -34.2856 | -15.4664
9.0892 | -34.286 | -15.4761
9.0953 | -34.2864 | -15.4858
9.1015 | -34.2867 | -15.4954
9.1076 | -34.2871 | -15.5048
9.1137 | -34.2875 | -15.5143
9.1198 | -34.2879 | -15.5239
9.1259 | -34.2883 | -15.5335
9.132 | -34.2886 | -15.5433
9.1381 | -34.289 | -15.5528
9.1442 | -34.2894 | -15.562
9.1503 | -34.2898 | -15.5718
9.1564 | -34.2902 | -15.5813
9.1625 | -34.2905 | -15.591
9.1686 | -34.291 | -15.6005
9.1747 | -34.2913 | -15.61
9.1808 | -34.2917 | -15.6195
9.1869 | -34.2921 | -15.6291
9.193 | -34.2925 | -15.6387
9.1991 | -34.2929 | -15.6483
9.2052 | -34.2933 | -15.6579
9.2113 | -34.2937 | -15.6675
9.2174 | -34.2941 | -15.677
9.2235 | -34.2945 | -15.6866
9.2296 | -34.2949 | -15.6962
9.2357 | -34.2952 | -15.7059
9.2419 | -34.2956 | -15.7154
9.248 | -34.2961 | -15.725
9.2541 | -34.2965 | -15.7346
9.2602 | -34.2969 | -15.7443
9.2663 | -34.2973 | -15.7536
9.2724 | -34.2977 | -15.7631
9.2785 | -34.298 | -15.7727
9.2846 | -34.2984 | -15.7821
9.2907 | -34.2988 | -15.7916
9.2968 | -34.2992 | -15.8013
9.3029 | -34.2996 | -15.811
9.309 | -34.3 | -15.8206
9.3151 | -34.3004 | -15.8301
9.3212 | -34.3008 | -15.8394
9.3273 | -34.3012 | -15.8489
9.3334 | -34.3016 | -15.8583
9.3395 | -34.302 | -15.8677
9.3456 | -34.3024 | -15.8771
9.3517 | -34.3027 | -15.8867
9.3578 | -34.3031 | -15.8962
9.3639 | -34.3035 | -15.9056
9.37 | -34.3039 | -15.9152
9.3761 | -34.3043 | -15.9249
9.3822 | -34.3047 | -15.9344
9.3884 | -34.3052 | -15.9437
9.3945 | -34.3055 | -15.9533
9.4006 | -34.306 | -15.9628
9.4067 | -34.3063 | -15.9724
9.4128 | -34.3067 | -15.982
9.4189 | -34.3072 | -15.9917
9.425 | -34.3076 | -16.0012
9.4311 | -34.308 | -16.0107
9.4372 | -34.3084 | -16.0202
9.4433 | -34.3088 | -16.0297
9.4494 | -34.3092 | -16.039
9.4555 | -34.3096 | -16.0486
9.4616 | -34.31 | -16.0581
9.4677 | -34.3104 | -16.0676
9.4738 | -34.3108 | -16.0771
9.4799 | -34.3112 | -16.0867
9.486 | -34.3116 | -16.0962
9.4921 | -34.312 | -16.1058
9.4982 | -34.3124 | -16.1151
9.5043 | -34.3128 | -16.1246
9.5104 | -34.3132 | -16.1339
9.5165 | -34.3136 | -16.1434
9.5226 | -34.314 | -16.1531
9.5288 | -34.3144 | -16.1624
9.5349 | -34.3148 | -16.1718
9.541 | -34.3152 | -16.1811
9.5471 | -34.3156 | -16.1905
9.5532 | -34.316 | -16.2001
9.5593 | -34.3165 | -16.2096
9.5654 | -34.3169 | -16.219
9.5715 | -34.3173 | -16.2285
9.5776 | -34.3177 | -16.2382
9.5837 | -34.318 | -16.2476
9.5898 | -34.3184 | -16.2572
9.5959 | -34.3188 | -16.2667
9.602 | -34.3192 | -16.2761
9.6081 | -34.3196 | -16.2855
9.6142 | -34.3201 | -16.2949
9.6203 | -34.3204 | -16.3045
9.6264 | -34.3208 | -16.3138
9.6325 | -34.3212 | -16.3234
9.6386 | -34.3216 | -16.3327
9.6447 | -34.322 | -16.3422
9.6508 | -34.3224 | -16.3515
9.6569 | -34.3228 | -16.3609
9.663 | -34.3233 | -16.3704
9.6691 | -34.3237 | -16.3799
9.6753 | -34.3241 | -16.3893
9.6814 | -34.3245 | -16.3986
9.6875 | -34.3249 | -16.4079
9.6936 | -34.3253 | -16.4172
9.6997 | -34.3257 | -16.4268
9.7058 | -34.3262 | -16.4361
9.7119 | -34.3266 | -16.4457
9.718 | -34.327 | -16.4553
9.7241 | -34.3274 | -16.4646
9.7302 | -34.3278 | -16.4741
9.7363 | -34.3282 | -16.4834
9.7424 | -34.3286 | -16.4926
9.7485 | -34.329 | -16.5021
9.7546 | -34.3294 | -16.5114
9.7607 | -34.3298 | -16.521
9.7668 | -34.3302 | -16.5305
9.7729 | -34.3306 | -16.5399
9.779 | -34.3311 | -16.549
9.7851 | -34.3315 | -16.5586
9.7912 | -34.3319 | -16.568
9.7973 | -34.3323 | -16.5775
9.8034 | -34.3328 | -16.5867
9.8095 | -34.3332 | -16.5963
9.8157 | -34.3336 | -16.6057
9.8218 | -34.334 | -16.6153
9.8279 | -34.3344 | -16.6246
9.834 | -34.3348 | -16.6341
9.8401 | -34.3352 | -16.6435
9.8462 | -34.3356 | -16.6527
9.8523 | -34.336 | -16.6621
9.8584 | -34.3364 | -16.6714
9.8645 | -34.3368 | -16.6809
9.8706 | -34.3372 | -16.6903
9.8767 | -34.3376 | -16.6997
9.8828 | -34.338 | -16.709
9.8889 | -34.3384 | -16.7182
9.895 | -34.3388 | -16.7276
9.9011 | -34.3392 | -16.737
9.9072 | -34.3396 | -16.7463
9.9133 | -34.34 | -16.7557
9.9194 | -34.3405 | -16.7651
9.9255 | -34.3409 | -16.7744
9.9316 | -34.3413 | -16.7838
9.9377 | -34.3417 | -16.7931
9.9438 | -34.3421 | -16.8026
9.9499 | -34.3425 | -16.8122
9.956 | -34.3429 | -16.8215
9.9622 | -34.3433 | -16.831
9.9683 | -34.3437 | -16.8404
9.9744 | -34.3441 | -16.8498
9.9805 | -34.3445 | -16.8592
9.9866 | -34.3449 | -16.8684
9.9927 | -34.3453 | -16.8777
9.9988 | -34.3457 | -16.8871
10.0049 | -34.3462 | -16.8966
10.011 | -34.3466 | -16.9061
10.0171 | -34.347 | -16.9154
10.0232 | -34.3474 | -16.9247
10.0293 | -34.3479 | -16.9341
10.0354 | -34.3483 | -16.9434
10.0415 | -34.3487 | -16.9528
10.0476 | -34.3491 | -16.9623
10.0537 | -34.3495 | -16.9717
10.0598 | -34.3499 | -16.9811
10.0659 | -34.3503 | -16.9904
10.072 | -34.3507 | -16.9996
10.0781 | -34.3511 | -17.0089
10.0842 | -34.3515 | -17.0182
10.0903 | -34.3519 | -17.0275
10.0964 | -34.3523 | -17.0367
10.1026 | -34.3527 | -17.046
10.1087 | -34.3531 | -17.0555
10.1148 | -34.3535 | -17.0648
10.1209 | -34.354 | -17.074
10.127 | -34.3544 | -17.0832
10.1331 | -34.3548 | -17.0925
10.1392 | -34.3552 | -17.1017
10.1453 | -34.3556 | -17.111
10.1514 | -34.356 | -17.1204
10.1575 | -34.3564 | -17.1299
10.1636 | -34.3568 | -17.139
10.1697 | -34.3572 | -17.1484
10.1758 | -34.3576 | -17.1574
10.1819 | -34.358 | -17.1668
10.188 | -34.3584 | -17.1762
10.1941 | -34.3588 | -17.1854
10.2002 | -34.3592 | -17.1947
10.2063 | -34.3596 | -17.2041
10.2124 | -34.36 | -17.2135
10.2185 | -34.3604 | -17.2227
10.2246 | -34.3609 | -17.2318
10.2307 | -34.3613 | -17.2411
10.2368 | -34.3617 | -17.2501
10.2429 | -34.3621 | -17.2594
10.2491 | -34.3625 | -17.2688
10.2552 | -34.3629 | -17.2782
10.2613 | -34.3633 | -17.2875
10.2674 | -34.3637 | -17.2967
10.2735 | -34.3641 | -17.3062
10.2796 | -34.3646 | -17.3153
10.2857 | -34.365 | -17.3247
10.2918 | -34.3654 | -17.334
10.2979 | -34.3658 | -17.3433
10.304 | -34.3662 | -17.3523
10.3101 | -34.3666 | -17.3616
10.3162 | -34.367 | -17.3709
10.3223 | -34.3674 | -17.38
10.3284 | -34.3678 | -17.3892
10.3345 | -34.3682 | -17.3983
10.3406 | -34.3686 | -17.4074
10.3467 | -34.369 | -17.4167
10.3528 | -34.3693 | -17.4258
10.3589 | -34.3697 | -17.435
10.365 | -34.3702 | -17.4445
10.3711 | -34.3705 | -17.4537
10.3772 | -34.371 | -17.4628
10.3833 | -34.3714 | -17.4721
10.3895 | -34.3718 | -17.4813
10.3956 | -34.3722 | -17.4905
10.4017 | -34.3727 | -17.4995
10.4078 | -34.373 | -17.5087
10.4139 | -34.3735 | -17.518
10.42 | -34.3739 | -17.5276
10.4261 | -34.3742 | -17.5367
10.4322 | -34.3747 | -17.5458
10.4383 | -34.375 | -17.555
10.4444 | -34.3754 | -17.5642
10.4505 | -34.3758 | -17.5735
10.4566 | -34.3762 | -17.5829
10.4627 | -34.3766 | -17.5921
10.4688 | -34.377 | -17.6013
10.4749 | -34.3774 | -17.6106
10.481 | -34.3778 | -17.6199
10.4871 | -34.3782 | -17.6289
10.4932 | -34.3786 | -17.638
10.4993 | -34.379 | -17.6473
10.5054 | -34.3794 | -17.6566
10.5115 | -34.3798 | -17.6658
10.5176 | -34.3802 | -17.6748
10.5237 | -34.3806 | -17.6842
10.5298 | -34.3809 | -17.6934
10.536 | -34.3813 | -17.7026
10.5421 | -34.3817 | -17.7118
10.5482 | -34.3821 | -17.7209
10.5543 | -34.3825 | -17.7298
10.5604 | -34.3829 | -17.7388
10.5665 | -34.3833 | -17.7479
10.5726 | -34.3836 | -17.7571
10.5787 | -34.384 | -17.7664
10.5848 | -34.3844 | -17.7753
10.5909 | -34.3848 | -17.7846
10.597 | -34.3852 | -17.794
10.6031 | -34.3856 | -17.8035
10.6092 | -34.386 | -17.8125
10.6153 | -34.3864 | -17.8219
10.6214 | -34.3868 | -17.8307
10.6275 | -34.3872 | -17.8399
10.6336 | -34.3876 | -17.849
10.6397 | -34.388 | -17.858
10.6458 | -34.3883 | -17.867
10.6519 | -34.3887 | -17.8762
10.658 | -34.3891 | -17.8851
10.6641 | -34.3895 | -17.8944
10.6702 | -34.3899 | -17.9036
10.6764 | -34.3903 | -17.9128
10.6825 | -34.3907 | -17.922
10.6886 | -34.3911 | -17.9311
10.6947 | -34.3915 | -17.9403
10.7008 | -34.3919 | -17.9494
10.7069 | -34.3923 | -17.9586
10.713 | -34.3927 | -17.9677
10.7191 | -34.3931 | -17.9768
10.7252 | -34.3935 | -17.9858
10.7313 | -34.3939 | -17.9949
10.7374 | -34.3942 | -18.0039
10.7435 | -34.3946 | -18.0131
10.7496 | -34.395 | -18.0221
10.7557 | -34.3954 | -18.0315
10.7618 | -34.3958 | -18.0403
10.7679 | -34.3962 | -18.0494
10.774 | -34.3966 | -18.0585
10.7801 | -34.397 | -18.0679
10.7862 | -34.3974 | -18.0771
10.7923 | -34.3978 | -18.0865
10.7984 | -34.3982 | -18.0954
10.8045 | -34.3986 | -18.1045
10.8106 | -34.399 | -18.1138
10.8168 | -34.3994 | -18.1229
10.8229 | -34.3998 | -18.1322
10.829 | -34.4002 | -18.1412
10.8351 | -34.4005 | -18.1502
10.8412 | -34.401 | -18.1592
10.8473 | -34.4014 | -18.1681
10.8534 | -34.4018 | -18.1773
10.8595 | -34.4022 | -18.1864
10.8656 | -34.4026 | -18.1955
10.8717 | -34.4029 | -18.2046
10.8778 | -34.4034 | -18.2136
10.8839 | -34.4037 | -18.2228
10.89 | -34.4041 | -18.232
10.8961 | -34.4046 | -18.2409
10.9022 | -34.4049 | -18.2501
10.9083 | -34.4053 | -18.2591
10.9144 | -34.4057 | -18.2682
10.9205 | -34.4061 | -18.2772
10.9266 | -34.4065 | -18.2862
10.9327 | -34.4069 | -18.2952
10.9388 | -34.4073 | -18.3038
10.9449 | -34.4077 | -18.3129
10.951 | -34.4081 | -18.3219
10.9571 | -34.4085 | -18.3311
10.9633 | -34.4088 | -18.3403
10.9694 | -34.4092 | -18.3495
10.9755 | -34.4096 | -18.3585
10.9816 | -34.41 | -18.3677
10.9877 | -34.4104 | -18.3766
10.9938 | -34.4108 | -18.3859
10.9999 | -34.4112 | -18.3948
11.006 | -34.4116 | -18.4037
11.0121 | -34.412 | -18.4129
11.0182 | -34.4123 | -18.4219
11.0243 | -34.4127 | -18.4311
11.0304 | -34.4131 | -18.4402
11.0365 | -34.4135 | -18.449
11.0426 | -34.4139 | -18.4578
11.0487 | -34.4144 | -18.467
11.0548 | -34.4148 | -18.4758
11.0609 | -34.4152 | -18.4849
11.067 | -34.4155 | -18.494
11.0731 | -34.4159 | -18.5028
11.0792 | -34.4163 | -18.5116
11.0853 | -34.4167 | -18.5206
11.0914 | -34.4171 | -18.5298
11.0975 | -34.4175 | -18.5388
11.1037 | -34.4179 | -18.5478
11.1098 | -34.4183 | -18.5569
11.1159 | -34.4187 | -18.5656
11.122 | -34.4191 | -18.5746
11.1281 | -34.4195 | -18.5835
11.1342 | -34.4199 | -18.5927
11.1403 | -34.4203 | -18.6018
11.1464 | -34.4207 | -18.6109
11.1525 | -34.421 | -18.6199
11.1586 | -34.4214 | -18.6291
11.1647 | -34.4218 | -18.638
11.1708 | -34.4222 | -18.6471
11.1769 | -34.4225 | -18.6563
11.183 | -34.423 | -18.6651
11.1891 | -34.4234 | -18.6741
11.1952 | -34.4238 | -18.683
11.2013 | -34.4242 | -18.6922
11.2074 | -34.4246 | -18.7012
11.2135 | -34.425 | -18.7103
11.2196 | -34.4253 | -18.7192
11.2257 | -34.4257 | -18.7284
11.2318 | -34.4262 | -18.7373
11.2379 | -34.4266 | -18.7464
11.244 | -34.427 | -18.7556
11.2502 | -34.4274 | -18.7645
11.2563 | -34.4278 | -18.7737
11.2624 | -34.4282 | -18.7826
11.2685 | -34.4286 | -18.7914
11.2746 | -34.429 | -18.8004
11.2807 | -34.4294 | -18.8097
11.2868 | -34.4298 | -18.8189
11.2929 | -34.4302 | -18.8276
11.299 | -34.4306 | -18.8366
11.3051 | -34.431 | -18.8459
11.3112 | -34.4314 | -18.855
11.3173 | -34.4318 | -18.8638
11.3234 | -34.4322 | -18.8726
11.3295 | -34.4326 | -18.8817
11.3356 | -34.433 | -18.8909
11.3417 | -34.4334 | -18.8998
11.3478 | -34.4338 | -18.9091
11.3539 | -34.4342 | -18.9181
11.36 | -34.4346 | -18.9271
11.3661 | -34.435 | -18.9359
11.3722 | -34.4354 | -18.9447
11.3783 | -34.4358 | -18.9537
11.3844 | -34.4362 | -18.9627
11.3906 | -34.4366 | -18.9717
11.3967 | -34.437 | -18.9806
11.4028 | -34.4374 | -18.9898
11.4089 | -34.4378 | -18.9987
11.415 | -34.4381 | -19.0076
11.4211 | -34.4385 | -19.0168
11.4272 | -34.4389 | -19.0259
11.4333 | -34.4393 | -19.0345
11.4394 | -34.4397 | -19.0435
11.4455 | -34.4401 | -19.0525
11.4516 | -34.4405 | -19.0616
11.4577 | -34.4409 | -19.0706
11.4638 | -34.4413 | -19.0794
11.4699 | -34.4417 | -19.0883
11.476 | -34.4421 | -19.0976
11.4821 | -34.4425 | -19.1066
11.4882 | -34.443 | -19.1158
11.4943 | -34.4433 | -19.1248
11.5004 | -34.4438 | -19.1337
11.5065 | -34.4442 | -19.1427
11.5126 | -34.4446 | -19.1516
11.5187 | -34.445 | -19.1606
11.5248 | -34.4454 | -19.1697
11.5309 | -34.4458 | -19.1787
11.5371 | -34.4462 | -19.1876
11.5432 | -34.4466 | -19.1966
11.5493 | -34.447 | -19.2057
11.5554 | -34.4474 | -19.2147
11.5615 | -34.4478 | -19.2238
11.5676 | -34.4482 | -19.2327
11.5737 | -34.4486 | -19.2417
11.5798 | -34.449 | -19.2504
11.5859 | -34.4494 | -19.2595
11.592 | -34.4498 | -19.2684
11.5981 | -34.4502 | -19.2773
11.6042 | -34.4507 | -19.2863
11.6103 | -34.451 | -19.2951
11.6164 | -34.4514 | -19.3039
11.6225 | -34.4518 | -19.3127
11.6286 | -34.4522 | -19.3215
11.6347 | -34.4526 | -19.3304
11.6408 | -34.453 | -19.3391
11.6469 | -34.4534 | -19.3481
11.653 | -34.4538 | -19.357
11.6591 | -34.4543 | -19.366
11.6652 | -34.4546 | -19.3748
11.6713 | -34.4551 | -19.3837
11.6775 | -34.4555 | -19.3924
11.6836 | -34.4559 | -19.4014
11.6897 | -34.4564 | -19.4103
11.6958 | -34.4568 | -19.4192
11.7019 | -34.4572 | -19.4283
11.708 | -34.4576 | -19.4372
11.7141 | -34.4581 | -19.4459
11.7202 | -34.4585 | -19.4548
11.7263 | -34.4589 | -19.4638
11.7324 | -34.4593 | -19.4728
11.7385 | -34.4597 | -19.4819
11.7446 | -34.46 | -19.4908
11.7507 | -34.4605 | -19.4997
11.7568 | -34.4609 | -19.5086
11.7629 | -34.4613 | -19.5177
11.769 | -34.4617 | -19.5265
11.7751 | -34.4621 | -19.5354
11.7812 | -34.4626 | -19.5441
11.7873 | -34.463 | -19.5531
11.7934 | -34.4634 | -19.562
11.7995 | -34.4638 | -19.5709
11.8056 | -34.4642 | -19.5797
11.8117 | -34.4646 | -19.5887
11.8178 | -34.4651 | -19.5974
11.824 | -34.4655 | -19.6062
11.8301 | -34.4659 | -19.615
11.8362 | -34.4663 | -19.6238
11.8423 | -34.4667 | -19.6327
11.8484 | -34.4671 | -19.6416
11.8545 | -34.4676 | -19.6504
11.8606 | -34.468 | -19.6591
11.8667 | -34.4684 | -19.6681
11.8728 | -34.4688 | -19.6768
11.8789 | -34.4692 | -19.6857
11.885 | -34.4696 | -19.6946
11.8911 | -34.4701 | -19.7034
11.8972 | -34.4705 | -19.7122
11.9033 | -34.4709 | -19.7209
11.9094 | -34.4713 | -19.7297
11.9155 | -34.4717 | -19.7385
11.9216 | -34.4722 | -19.7472
11.9277 | -34.4726 | -19.756
11.9338 | -34.473 | -19.7648
11.9399 | -34.4734 | -19.7738
11.946 | -34.4738 | -19.7825
11.9521 | -34.4742 | -19.7912
11.9582 | -34.4747 | -19.8
11.9644 | -34.4751 | -19.8087
11.9705 | -34.4755 | -19.8175
11.9766 | -34.476 | -19.8264
11.9827 | -34.4764 | -19.8352
11.9888 | -34.4769 | -19.8439
11.9949 | -34.4773 | -19.8525
12.001 | -34.4777 | -19.8614
12.0071 | -34.4782 | -19.8702
12.0132 | -34.4786 | -19.8788
12.0193 | -34.479 | -19.8877
12.0254 | -34.4794 | -19.8964
12.0315 | -34.4798 | -19.905
12.0376 | -34.4803 | -19.9137
12.0437 | -34.4807 | -19.9225
12.0498 | -34.4811 | -19.9313
12.0559 | -34.4815 | -19.9399
12.062 | -34.482 | -19.9485
12.0681 | -34.4824 | -19.9573
12.0742 | -34.4828 | -19.966
12.0803 | -34.4833 | -19.9749
12.0864 | -34.4837 | -19.9839
12.0925 | -34.4841 | -19.9926
12.0986 | -34.4846 | -20.0013
12.1047 | -34.485 | -20.0101
12.1109 | -34.4854 | -20.0188
12.117 | -34.4859 | -20.0275
12.1231 | -34.4863 | -20.0362
12.1292 | -34.4867 | -20.0448
12.1353 | -34.4872 | -20.0535
12.1414 | -34.4876 | -20.0623
12.1475 | -34.488 | -20.0712
12.1536 | -34.4884 | -20.0799
12.1597 | -34.4889 | -20.0887
12.1658 | -34.4893 | -20.0974
12.1719 | -34.4898 | -20.1059
12.178 | -34.4902 | -20.1147
12.1841 | -34.4906 | -20.1235
12.1902 | -34.491 | -20.1322
12.1963 | -34.4915 | -20.141
12.2024 | -34.4919 | -20.1496
12.2085 | -34.4924 | -20.1583
12.2146 | -34.4928 | -20.1671
12.2207 | -34.4933 | -20.1758
12.2268 | -34.4937 | -20.1845
12.2329 | -34.4941 | -20.1932
12.239 | -34.4946 | -20.2015
12.2451 | -34.495 | -20.2102
12.2513 | -34.4955 | -20.2191
12.2574 | -34.4959 | -20.2278
12.2635 | -34.4963 | -20.2365
12.2696 | -34.4968 | -20.2451
12.2757 | -34.4973 | -20.2536
12.2818 | -34.4977 | -20.2624
12.2879 | -34.4981 | -20.2711
12.294 | -34.4986 | -20.2797
12.3001 | -34.499 | -20.2884
12.3062 | -34.4995 | -20.2971
12.3123 | -34.4999 | -20.3054
12.3184 | -34.5004 | -20.3141
12.3245 | -34.5008 | -20.3225
12.3306 | -34.5012 | -20.3313
12.3367 | -34.5017 | -20.3399
12.3428 | -34.5021 | -20.3485
12.3489 | -34.5026 | -20.3569
12.355 | -34.5031 | -20.3657
12.3611 | -34.5035 | -20.3745
12.3672 | -34.5039 | -20.383
12.3733 | -34.5044 | -20.3917
12.3794 | -34.5049 | -20.4
12.3855 | -34.5053 | -20.4086
12.3916 | -34.5057 | -20.4172
12.3978 | -34.5062 | -20.4261
12.4039 | -34.5067 | -20.4347
12.41 | -34.5071 | -20.4434
12.4161 | -34.5076 | -20.4517
12.4222 | -34.508 | -20.4602
12.4283 | -34.5085 | -20.4687
12.4344 | -34.5089 | -20.4773
12.4405 | -34.5094 | -20.4861
12.4466 | -34.5098 | -20.4946
12.4527 | -34.5103 | -20.5031
12.4588 | -34.5107 | -20.5117
12.4649 | -34.5112 | -20.5203
12.471 | -34.5116 | -20.529
12.4771 | -34.5121 | -20.5375
12.4832 | -34.5125 | -20.5462
12.4893 | -34.513 | -20.5545
12.4954 | -34.5135 | -20.563
12.5015 | -34.5139 | -20.5716
12.5076 | -34.5143 | -20.5802
12.5137 | -34.5148 | -20.5889
12.5198 | -34.5153 | -20.5974
12.5259 | -34.5157 | -20.6061
12.532 | -34.5161 | -20.6146
12.5382 | -34.5166 | -20.6232
12.5443 | -34.5171 | -20.6317
12.5504 | -34.5175 | -20.6404
12.5565 | -34.518 | -20.6488
12.5626 | -34.5185 | -20.6575
12.5687 | -34.5189 | -20.6661
12.5748 | -34.5194 | -20.6747
12.5809 | -34.5198 | -20.6832
12.587 | -34.5203 | -20.6919
12.5931 | -34.5208 | -20.7003
12.5992 | -34.5213 | -20.7089
12.6053 | -34.5217 | -20.7174
12.6114 | -34.5222 | -20.7259
12.6175 | -34.5226 | -20.7346
12.6236 | -34.5231 | -20.7432
12.6297 | -34.5236 | -20.7518
12.6358 | -34.524 | -20.7601
12.6419 | -34.5244 | -20.7687
12.648 | -34.5249 | -20.7773
12.6541 | -34.5254 | -20.7857
12.6602 | -34.5259 | -20.7937
12.6663 | -34.5263 | -20.8022
12.6724 | -34.5268 | -20.8107
12.6785 | -34.5273 | -20.8193
12.6847 | -34.5277 | -20.8277
12.6908 | -34.5282 | -20.8362
12.6969 | -34.5287 | -20.8445
12.703 | -34.5291 | -20.8531
12.7091 | -34.5296 | -20.8617
12.7152 | -34.53 | -20.87
12.7213 | -34.5305 | -20.8787
12.7274 | -34.531 | -20.8872
12.7335 | -34.5315 | -20.8954
12.7396 | -34.5319 | -20.9038
12.7457 | -34.5324 | -20.9125
12.7518 | -34.5328 | -20.9209
12.7579 | -34.5333 | -20.9296
12.764 | -34.5337 | -20.9381
12.7701 | -34.5342 | -20.9466
12.7762 | -34.5347 | -20.955
12.7823 | -34.5352 | -20.9634
12.7884 | -34.5357 | -20.9719
12.7945 | -34.5361 | -20.9806
12.8006 | -34.5366 | -20.989
12.8067 | -34.5371 | -20.9973
12.8128 | -34.5375 | -21.0057
12.8189 | -34.538 | -21.0141
12.8251 | -34.5385 | -21.0227
12.8312 | -34.5389 | -21.0313
12.8373 | -34.5394 | -21.0398
12.8434 | -34.5398 | -21.0484
12.8495 | -34.5403 | -21.0567
12.8556 | -34.5407 | -21.0652
12.8617 | -34.5412 | -21.0738
12.8678 | -34.5416 | -21.0824
12.8739 | -34.5421 | -21.0907
12.88 | -34.5426 | -21.0991
12.8861 | -34.543 | -21.1074
12.8922 | -34.5435 | -21.1159
12.8983 | -34.5439 | -21.1245
12.9044 | -34.5444 | -21.1329
12.9105 | -34.5449 | -21.1411
12.9166 | -34.5453 | -21.1494
12.9227 | -34.5458 | -21.1578
12.9288 | -34.5462 | -21.1664
12.9349 | -34.5467 | -21.1748
12.941 | -34.5472 | -21.1831
12.9471 | -34.5477 | -21.1916
12.9532 | -34.5481 | -21.2001
12.9593 | -34.5486 | -21.2088
12.9654 | -34.5491 | -21.2172
12.9716 | -34.5495 | -21.2258
12.9777 | -34.55 | -21.2343
12.9838 | -34.5505 | -21.2428
12.9899 | -34.5509 | -21.2514
12.996 | -34.5514 | -21.2598
13.0021 | -34.5519 | -21.2681
13.0082 | -34.5524 | -21.2766
13.0143 | -34.5529 | -21.2847
13.0204 | -34.5533 | -21.293
13.0265 | -34.5538 | -21.3015
13.0326 | -34.5543 | -21.31
13.0387 | -34.5548 | -21.3185
13.0448 | -34.5552 | -21.327
13.0509 | -34.5557 | -21.3352
13.057 | -34.5562 | -21.3435
13.0631 | -34.5567 | -21.3519
13.0692 | -34.5571 | -21.3603
13.0753 | -34.5576 | -21.3689
13.0814 | -34.5581 | -21.3773
13.0875 | -34.5585 | -21.3856
13.0936 | -34.559 | -21.3941
13.0997 | -34.5594 | -21.4024
13.1058 | -34.5599 | -21.4108
13.112 | -34.5604 | -21.4192
13.1181 | -34.5609 | -21.4275
13.1242 | -34.5613 | -21.4359
13.1303 | -34.5618 | -21.4444
13.1364 | -34.5623 | -21.4527
13.1425 | -34.5627 | -21.461
13.1486 | -34.5632 | -21.4693
13.1547 | -34.5637 | -21.4776
13.1608 | -34.5642 | -21.486
13.1669 | -34.5646 | -21.4942
13.173 | -34.5651 | -21.5027
13.1791 | -34.5656 | -21.5109
13.1852 | -34.566 | -21.5193
13.1913 | -34.5665 | -21.5274
13.1974 | -34.567 | -21.5359
13.2035 | -34.5674 | -21.5442
13.2096 | -34.5679 | -21.5525
13.2157 | -34.5684 | -21.5608
13.2218 | -34.5689 | -21.5691
13.2279 | -34.5693 | -21.5773
13.234 | -34.5698 | -21.5855
13.2401 | -34.5703 | -21.5939
13.2462 | -34.5707 | -21.6024
13.2524 | -34.5712 | -21.6109
13.2585 | -34.5717 | -21.6192
13.2646 | -34.5721 | -21.6275
13.2707 | -34.5726 | -21.6359
13.2768 | -34.5731 | -21.6443
13.2829 | -34.5735 | -21.6528
13.289 | -34.574 | -21.661
13.2951 | -34.5745 | -21.669
13.3012 | -34.575 | -21.6774
13.3073 | -34.5754 | -21.6858
13.3134 | -34.5759 | -21.6942
13.3195 | -34.5764 | -21.7024
13.3256 | -34.5768 | -21.7106
13.3317 | -34.5773 | -21.7188
13.3378 | -34.5777 | -21.7273
13.3439 | -34.5782 | -21.7358
13.35 | -34.5787 | -21.7442
13.3561 | -34.5791 | -21.7524
13.3622 | -34.5796 | -21.7608
13.3683 | -34.5801 | -21.7691
13.3744 | -34.5805 | -21.7774
13.3805 | -34.581 | -21.7857
13.3866 | -34.5814 | -21.7939
13.3927 | -34.5819 | -21.8022
13.3989 | -34.5824 | -21.8105
13.405 | -34.5828 | -21.8188
13.4111 | -34.5833 | -21.8271
13.4172 | -34.5838 | -21.8356
13.4233 | -34.5842 | -21.8441
13.4294 | -34.5847 | -21.8526
13.4355 | -34.5851 | -21.8608
13.4416 | -34.5856 | -21.8691
13.4477 | -34.5861 | -21.8773
13.4538 | -34.5865 | -21.8857
13.4599 | -34.587 | -21.8941
13.466 | -34.5875 | -21.9025
13.4721 | -34.588 | -21.9108
13.4782 | -34.5884 | -21.9192
13.4843 | -34.5889 | -21.9275
13.4904 | -34.5893 | -21.9358
13.4965 | -34.5898 | -21.9442
13.5026 | -34.5903 | -21.9528
13.5087 | -34.5907 | -21.9609
13.5148 | -34.5912 | -21.9694
13.5209 | -34.5916 | -21.9779
13.527 | -34.5921 | -21.9862
13.5331 | -34.5926 | -21.9945
13.5393 | -34.593 | -22.0026
13.5454 | -34.5935 | -22.0108
13.5515 | -34.594 | -22.0193
13.5576 | -34.5944 | -22.0278
13.5637 | -34.5949 | -22.0362
13.5698 | -34.5954 | -22.0446
13.5759 | -34.5959 | -22.0526
13.582 | -34.5963 | -22.0609
13.5881 | -34.5968 | -22.0692
13.5942 | -34.5972 | -22.0777
13.6003 | -34.5977 | -22.086
13.6064 | -34.5982 | -22.0944
13.6125 | -34.5986 | -22.1029
13.6186 | -34.5991 | -22.1111
13.6247 | -34.5996 | -22.1195
13.6308 | -34.6 | -22.1278
13.6369 | -34.6005 | -22.1362
13.643 | -34.601 | -22.1447
13.6491 | -34.6014 | -22.1531
13.6552 | -34.6019 | -22.1613
13.6613 | -34.6023 | -22.1696
13.6674 | -34.6028 | -22.1779
13.6735 | -34.6032 | -22.1862
13.6796 | -34.6037 | -22.1945
13.6858 | -34.6041 | -22.2028
13.6919 | -34.6046 | -22.211
13.698 | -34.6051 | -22.2194
13.7041 | -34.6055 | -22.2278
13.7102 | -34.606 | -22.2362
13.7163 | -34.6065 | -22.2445
13.7224 | -34.6069 | -22.2528
13.7285 | -34.6074 | -22.261
13.7346 | -34.6079 | -22.2694
13.7407 | -34.6084 | -22.2777
13.7468 | -34.6088 | -22.2861
13.7529 | -34.6093 | -22.2943
13.759 | -34.6097 | -22.3028
13.7651 | -34.6102 | -22.3112
13.7712 | -34.6106 | -22.3196
13.7773 | -34.6111 | -22.3279
13.7834 | -34.6116 | -22.3362
13.7895 | -34.612 | -22.3445
13.7956 | -34.6125 | -22.353
13.8017 | -34.6129 | -22.3612
13.8078 | -34.6134 | -22.3695
13.8139 | -34.6138 | -22.3777
13.82 | -34.6143 | -22.3859
13.8262 | -34.6147 | -22.3943
13.8323 | -34.6152 | -22.4027
13.8384 | -34.6157 | -22.411
13.8445 | -34.6161 | -22.4193
13.8506 | -34.6166 | -22.4276
13.8567 | -34.617 | -22.4362
13.8628 | -34.6175 | -22.4445
13.8689 | -34.618 | -22.4528
13.875 | -34.6184 | -22.4611
13.8811 | -34.6189 | -22.4694
13.8872 | -34.6193 | -22.4776
13.8933 | -34.6198 | -22.4859
13.8994 | -34.6202 | -22.4943
13.9055 | -34.6207 | -22.5028
13.9116 | -34.6211 | -22.5111
13.9177 | -34.6216 | -22.5194
13.9238 | -34.622 | -22.5277
13.9299 | -34.6225 | -22.5361
13.936 | -34.6229 | -22.5445
13.9421 | -34.6234 | -22.5527
13.9482 | -34.6238 | -22.5609
13.9543 | -34.6242 | -22.5693
13.9604 | -34.6247 | -22.5775
13.9665 | -34.6252 | -22.5857
13.9727 | -34.6256 | -22.5943
13.9788 | -34.6261 | -22.6025
13.9849 | -34.6265 | -22.6109
13.991 | -34.627 | -22.6192
13.9971 | -34.6274 | -22.6274
14.0032 | -34.6279 | -22.6357
14.0093 | -34.6283 | -22.6439
14.0154 | -34.6288 | -22.6522
14.0215 | -34.6292 | -22.6604
14.0276 | -34.6297 | -22.6687
14.0337 | -34.6301 | -22.677
14.0398 | -34.6306 | -22.6853
14.0459 | -34.6311 | -22.6936
14.052 | -34.6315 | -22.702
14.0581 | -34.6319 | -22.7104
14.0642 | -34.6324 | -22.7187
14.0703 | -34.6328 | -22.7271
14.0764 | -34.6333 | -22.7352
14.0825 | -34.6337 | -22.7436
14.0886 | -34.6342 | -22.7519
14.0947 | -34.6346 | -22.7601
14.1008 | -34.6351 | -22.7684
14.1069 | -34.6355 | -22.7765
14.1131 | -34.6359 | -22.7849
14.1192 | -34.6364 | -22.7932
14.1253 | -34.6368 | -22.8013
14.1314 | -34.6373 | -22.8097
14.1375 | -34.6378 | -22.8179
14.1436 | -34.6382 | -22.826
14.1497 | -34.6387 | -22.8343
14.1558 | -34.6391 | -22.8426
14.1619 | -34.6396 | -22.8508
14.168 | -34.64 | -22.8591
14.1741 | -34.6405 | -22.8674
14.1802 | -34.6409 | -22.8757
14.1863 | -34.6414 | -22.8841
14.1924 | -34.6419 | -22.8923
14.1985 | -34.6423 | -22.9006
14.2046 | -34.6428 | -22.909
14.2107 | -34.6432 | -22.9171
14.2168 | -34.6437 | -22.9255
14.2229 | -34.6441 | -22.9337
14.229 | -34.6446 | -22.9421
14.2351 | -34.645 | -22.9503
14.2412 | -34.6454 | -22.9586
14.2473 | -34.6459 | -22.9667
14.2534 | -34.6463 | -22.9749
14.2596 | -34.6468 | -22.9831
14.2657 | -34.6472 | -22.9915
14.2718 | -34.6477 | -22.9998
14.2779 | -34.6481 | -23.0081
14.284 | -34.6486 | -23.0164
14.2901 | -34.649 | -23.0247
14.2962 | -34.6495 | -23.033
14.3023 | -34.6499 | -23.0413
14.3084 | -34.6504 | -23.0496
14.3145 | -34.6508 | -23.0579
14.3206 | -34.6513 | -23.0664
14.3267 | -34.6517 | -23.0746
14.3328 | -34.6522 | -23.0829
14.3389 | -34.6526 | -23.0912
14.345 | -34.653 | -23.0994
14.3511 | -34.6535 | -23.1077
14.3572 | -34.6539 | -23.1161
14.3633 | -34.6544 | -23.1245
14.3694 | -34.6548 | -23.1326
14.3755 | -34.6553 | -23.141
14.3816 | -34.6557 | -23.1495
14.3877 | -34.6562 | -23.1576
14.3938 | -34.6567 | -23.1661
14.4 | -34.6571 | -23.1742
14.4061 | -34.6575 | -23.1824
14.4122 | -34.658 | -23.1908
14.4183 | -34.6584 | -23.1989
14.4244 | -34.6589 | -23.2071
14.4305 | -34.6594 | -23.2154
14.4366 | -34.6598 | -23.2237
14.4427 | -34.6603 | -23.2321
14.4488 | -34.6607 | -23.2403
14.4549 | -34.6612 | -23.2485
14.461 | -34.6616 | -23.2567
14.4671 | -34.6621 | -23.2651
14.4732 | -34.6625 | -23.2734
14.4793 | -34.663 | -23.2817
14.4854 | -34.6635 | -23.29
14.4915 | -34.6639 | -23.2981
14.4976 | -34.6644 | -23.3062
14.5037 | -34.6648 | -23.3144
14.5098 | -34.6653 | -23.3226
14.5159 | -34.6657 | -23.3309
14.522 | -34.6661 | -23.3392
14.5281 | -34.6666 | -23.3474
14.5342 | -34.667 | -23.3559
14.5403 | -34.6675 | -23.364
14.5465 | -34.6679 | -23.3722
14.5526 | -34.6684 | -23.3807
14.5587 | -34.6688 | -23.389
14.5648 | -34.6693 | -23.3972
14.5709 | -34.6698 | -23.4054
14.577 | -34.6702 | -23.4135
14.5831 | -34.6706 | -23.4218
14.5892 | -34.6711 | -23.4301
14.5953 | -34.6715 | -23.4383
14.6014 | -34.672 | -23.4465
14.6075 | -34.6724 | -23.4546
14.6136 | -34.6728 | -23.4631
14.6197 | -34.6733 | -23.4714
14.6258 | -34.6737 | -23.4797
14.6319 | -34.6742 | -23.4879
14.638 | -34.6746 | -23.4961
14.6441 | -34.675 | -23.5044
14.6502 | -34.6755 | -23.5128
14.6563 | -34.676 | -23.5212
14.6624 | -34.6764 | -23.5295
14.6685 | -34.6769 | -23.5376
14.6746 | -34.6773 | -23.5459
14.6807 | -34.6778 | -23.554
14.6869 | -34.6782 | -23.5624
14.693 | -34.6787 | -23.5705
14.6991 | -34.6791 | -23.5788
14.7052 | -34.6796 | -23.5869
14.7113 | -34.6801 | -23.5952
14.7174 | -34.6806 | -23.6035
14.7235 | -34.681 | -23.6117
14.7296 | -34.6814 | -23.6198
14.7357 | -34.6819 | -23.6279
14.7418 | -34.6824 | -23.6362
14.7479 | -34.6828 | -23.6443
14.754 | -34.6833 | -23.6527
14.7601 | -34.6837 | -23.6608
14.7662 | -34.6842 | -23.6689
14.7723 | -34.6846 | -23.6768
14.7784 | -34.685 | -23.6851
14.7845 | -34.6855 | -23.6934
14.7906 | -34.6859 | -23.7017
14.7967 | -34.6864 | -23.7098
14.8028 | -34.6868 | -23.718
14.8089 | -34.6873 | -23.7261
14.815 | -34.6877 | -23.7342
14.8211 | -34.6882 | -23.7425
14.8272 | -34.6886 | -23.7508
14.8334 | -34.6891 | -23.7589
14.8395 | -34.6896 | -23.7669
14.8456 | -34.69 | -23.775
14.8517 | -34.6904 | -23.7833
14.8578 | -34.6909 | -23.7916
14.8639 | -34.6914 | -23.7997
14.87 | -34.6918 | -23.8078
14.8761 | -34.6923 | -23.8159
14.8822 | -34.6928 | -23.824
14.8883 | -34.6932 | -23.8322
14.8944 | -34.6937 | -23.8402
14.9005 | -34.6941 | -23.8485
14.9066 | -34.6946 | -23.8567
14.9127 | -34.6951 | -23.8649
14.9188 | -34.6955 | -23.873
14.9249 | -34.6959 | -23.8811
14.931 | -34.6964 | -23.8894
14.9371 | -34.6969 | -23.8975
14.9432 | -34.6973 | -23.9057
14.9493 | -34.6978 | -23.9139
14.9554 | -34.6982 | -23.9222
14.9615 | -34.6987 | -23.9304
14.9676 | -34.6992 | -23.9386
14.9738 | -34.6996 | -23.9467
14.9799 | -34.7001 | -23.9549
14.986 | -34.7006 | -23.963
14.9921 | -34.701 | -23.9713
14.9982 | -34.7014 | -23.9795
15.0043 | -34.7019 | -23.9876
15.0104 | -34.7023 | -23.9959
15.0165 | -34.7028 | -24.0042
15.0226 | -34.7033 | -24.0121
15.0287 | -34.7037 | -24.0203
15.0348 | -34.7042 | -24.0285
15.0409 | -34.7046 | -24.0366
15.047 | -34.7051 | -24.0449
15.0531 | -34.7056 | -24.053
15.0592 | -34.706 | -24.0612
15.0653 | -34.7065 | -24.0693
15.0714 | -34.7069 | -24.0772
15.0775 | -34.7074 | -24.0853
15.0836 | -34.7078 | -24.0935
15.0897 | -34.7083 | -24.1014
15.0958 | -34.7088 | -24.1096
15.1019 | -34.7092 | -24.1179
15.108 | -34.7097 | -24.1261
15.1141 | -34.7101 | -24.1343
15.1203 | -34.7106 | -24.1425
15.1264 | -34.7111 | -24.1506
15.1325 | -34.7116 | -24.1588
15.1386 | -34.712 | -24.167
15.1447 | -34.7125 | -24.1751
15.1508 | -34.713 | -24.1834
15.1569 | -34.7134 | -24.1915
15.163 | -34.7139 | -24.1993
15.1691 | -34.7144 | -24.2074
15.1752 | -34.7148 | -24.2155
15.1813 | -34.7153 | -24.2234
15.1874 | -34.7157 | -24.2315
15.1935 | -34.7162 | -24.2394
15.1996 | -34.7167 | -24.2475
15.2057 | -34.7171 | -24.2555
15.2118 | -34.7176 | -24.2637
15.2179 | -34.718 | -24.2718
15.224 | -34.7185 | -24.2799
15.2301 | -34.719 | -24.2876
15.2362 | -34.7194 | -24.2958
15.2423 | -34.7199 | -24.3039
15.2484 | -34.7203 | -24.3121
15.2545 | -34.7208 | -24.3203
15.2607 | -34.7212 | -24.3285
15.2668 | -34.7217 | -24.3363
15.2729 | -34.7222 | -24.3443
15.279 | -34.7227 | -24.3525
15.2851 | -34.7232 | -24.3604
15.2912 | -34.7236 | -24.3686
15.2973 | -34.7241 | -24.3768
15.3034 | -34.7245 | -24.3847
15.3095 | -34.725 | -24.3929
15.3156 | -34.7255 | -24.4009
15.3217 | -34.7259 | -24.4087
15.3278 | -34.7264 | -24.4167
15.3339 | -34.7269 | -24.4247
15.34 | -34.7274 | -24.4326
15.3461 | -34.7278 | -24.4407
15.3522 | -34.7282 | -24.4488
15.3583 | -34.7287 | -24.457
15.3644 | -34.7292 | -24.4649
15.3705 | -34.7297 | -24.4727
15.3766 | -34.7301 | -24.4807
15.3827 | -34.7306 | -24.4886
15.3888 | -34.7311 | -24.4968
15.3949 | -34.7315 | -24.5048
15.401 | -34.732 | -24.5129
15.4072 | -34.7325 | -24.5208
15.4133 | -34.7329 | -24.5288
15.4194 | -34.7334 | -24.5369
15.4255 | -34.7339 | -24.5449
15.4316 | -34.7343 | -24.553
15.4377 | -34.7348 | -24.5612
15.4438 | -34.7353 | -24.5691
15.4499 | -34.7358 | -24.577
15.456 | -34.7362 | -24.585
15.4621 | -34.7367 | -24.593
15.4682 | -34.7372 | -24.6011
15.4743 | -34.7377 | -24.6089
15.4804 | -34.7381 | -24.6169
15.4865 | -34.7385 | -24.625
15.4926 | -34.739 | -24.633
15.4987 | -34.7395 | -24.641
15.5048 | -34.74 | -24.6491
15.5109 | -34.7404 | -24.6571
15.517 | -34.7409 | -24.6651
15.5231 | -34.7414 | -24.6732
15.5292 | -34.7419 | -24.6811
15.5353 | -34.7423 | -24.6893
15.5414 | -34.7428 | -24.6973
15.5476 | -34.7433 | -24.7051
15.5537 | -34.7438 | -24.7132
15.5598 | -34.7442 | -24.7214
15.5659 | -34.7447 | -24.7293
15.572 | -34.7452 | -24.7371
15.5781 | -34.7456 | -24.7451
15.5842 | -34.7461 | -24.7529
15.5903 | -34.7466 | -24.7608
15.5964 | -34.7471 | -24.7688
15.6025 | -34.7475 | -24.7768
15.6086 | -34.748 | -24.7848
15.6147 | -34.7485 | -24.7928
15.6208 | -34.749 | -24.8005
15.6269 | -34.7495 | -24.8083
15.633 | -34.75 | -24.8163
15.6391 | -34.7504 | -24.8243
15.6452 | -34.7509 | -24.8323
15.6513 | -34.7514 | -24.8399
15.6574 | -34.7519 | -24.8479
15.6635 | -34.7524 | -24.8558
15.6696 | -34.7528 | -24.8638
15.6757 | -34.7533 | -24.8718
15.6818 | -34.7538 | -24.8799
15.688 | -34.7542 | -24.8877
15.6941 | -34.7547 | -24.8957
15.7002 | -34.7551 | -24.9037
15.7063 | -34.7556 | -24.9116
15.7124 | -34.7561 | -24.9197
15.7185 | -34.7565 | -24.9276
15.7246 | -34.757 | -24.9354
15.7307 | -34.7575 | -24.9433
15.7368 | -34.758 | -24.9514
15.7429 | -34.7585 | -24.9594
15.749 | -34.759 | -24.9674
15.7551 | -34.7595 | -24.9754
15.7612 | -34.76 | -24.9832
15.7673 | -34.7604 | -24.9911
15.7734 | -34.7609 | -24.999
15.7795 | -34.7614 | -25.007
15.7856 | -34.7618 | -25.0148
15.7917 | -34.7623 | -25.0224
15.7978 | -34.7628 | -25.0305
15.8039 | -34.7633 | -25.0385
15.81 | -34.7638 | -25.0464
15.8161 | -34.7642 | -25.0543
15.8222 | -34.7647 | -25.0621
15.8283 | -34.7652 | -25.0698
15.8345 | -34.7657 | -25.0777
15.8406 | -34.7662 | -25.0857
15.8467 | -34.7667 | -25.0936
15.8528 | -34.7671 | -25.1014
15.8589 | -34.7676 | -25.1093
15.865 | -34.7682 | -25.1169
15.8711 | -34.7686 | -25.1249
15.8772 | -34.7691 | -25.1327
15.8833 | -34.7696 | -25.1405
15.8894 | -34.7701 | -25.1485
15.8955 | -34.7705 | -25.1563
15.9016 | -34.771 | -25.1641
15.9077 | -34.7715 | -25.1721
15.9138 | -34.7719 | -25.1801
15.9199 | -34.7724 | -25.1877
15.926 | -34.7729 | -25.1955
15.9321 | -34.7734 | -25.2032
15.9382 | -34.7739 | -25.2108
15.9443 | -34.7744 | -25.2187
15.9504 | -34.7749 | -25.2267
15.9565 | -34.7753 | -25.2345
15.9626 | -34.7758 | -25.2425
15.9687 | -34.7763 | -25.25
15.9749 | -34.7767 | -25.2576
15.981 | -34.7772 | -25.2655
15.9871 | -34.7777 | -25.2735
15.9932 | -34.7781 | -25.2813
15.9993 | -34.7786 | -25.2892
16.0054 | -34.7791 | -25.2968
16.0115 | -34.7796 | -25.3047
16.0176 | -34.7801 | -25.3126
16.0237 | -34.7806 | -25.3205
16.0298 | -34.781 | -25.3282
16.0359 | -34.7815 | -25.336
16.042 | -34.782 | -25.3435
16.0481 | -34.7825 | -25.3514
16.0542 | -34.783 | -25.3593
16.0603 | -34.7835 | -25.3671
16.0664 | -34.7839 | -25.3749
16.0725 | -34.7845 | -25.3824
16.0786 | -34.785 | -25.3901
16.0847 | -34.7854 | -25.398
16.0908 | -34.7859 | -25.4058
16.0969 | -34.7864 | -25.4138
16.103 | -34.7869 | -25.4216
16.1091 | -34.7874 | -25.4291
16.1152 | -34.7879 | -25.4368
16.1214 | -34.7883 | -25.4447
16.1275 | -34.7888 | -25.4525
16.1336 | -34.7893 | -25.4602
16.1397 | -34.7898 | -25.4679
16.1458 | -34.7903 | -25.4756
16.1519 | -34.7908 | -25.4836
16.158 | -34.7913 | -25.4914
16.1641 | -34.7918 | -25.4994
16.1702 | -34.7922 | -25.5072
16.1763 | -34.7927 | -25.5151
16.1824 | -34.7932 | -25.5227
16.1885 | -34.7937 | -25.5304
16.1946 | -34.7942 | -25.5381
16.2007 | -34.7947 | -25.5458
16.2068 | -34.7952 | -25.5535
16.2129 | -34.7957 | -25.5614
16.219 | -34.7962 | -25.5689
16.2251 | -34.7967 | -25.5766
16.2312 | -34.7972 | -25.5843
16.2373 | -34.7977 | -25.5919
16.2434 | -34.7982 | -25.5998
16.2495 | -34.7987 | -25.6074
16.2556 | -34.7992 | -25.6152
16.2618 | -34.7997 | -25.6229
16.2679 | -34.8002 | -25.6307
16.274 | -34.8006 | -25.6383
16.2801 | -34.8011 | -25.6461
16.2862 | -34.8016 | -25.6537
16.2923 | -34.8021 | -25.6615
16.2984 | -34.8026 | -25.6693
16.3045 | -34.8031 | -25.6772
16.3106 | -34.8036 | -25.6849
16.3167 | -34.8041 | -25.6926
16.3228 | -34.8046 | -25.7001
16.3289 | -34.805 | -25.7079
16.335 | -34.8055 | -25.7156
16.3411 | -34.806 | -25.7233
16.3472 | -34.8065 | -25.7311
16.3533 | -34.807 | -25.7388
16.3594 | -34.8075 | -25.7464
16.3655 | -34.8079 | -25.754
16.3716 | -34.8084 | -25.7616
16.3777 | -34.8089 | -25.7694
16.3838 | -34.8094 | -25.7772
16.3899 | -34.8099 | -25.7847
16.396 | -34.8104 | -25.7924
16.4021 | -34.8109 | -25.7999
16.4083 | -34.8113 | -25.8076
16.4144 | -34.8118 | -25.8154
16.4205 | -34.8123 | -25.8231
16.4266 | -34.8128 | -25.8305
16.4327 | -34.8133 | -25.8383
16.4388 | -34.8138 | -25.8461
16.4449 | -34.8142 | -25.8537
16.451 | -34.8147 | -25.8616
16.4571 | -34.8152 | -25.8694
16.4632 | -34.8157 | -25.8766
16.4693 | -34.8161 | -25.8843
16.4754 | -34.8166 | -25.8921
16.4815 | -34.8171 | -25.8998
16.4876 | -34.8176 | -25.9075
16.4937 | -34.8181 | -25.9152
16.4998 | -34.8186 | -25.9228
16.5059 | -34.8191 | -25.9306
16.512 | -34.8196 | -25.9385
16.5181 | -34.82 | -25.9464
16.5242 | -34.8205 | -25.9543
16.5303 | -34.821 | -25.9618
16.5364 | -34.8215 | -25.9695
16.5425 | -34.822 | -25.9773
16.5487 | -34.8225 | -25.985
16.5548 | -34.823 | -25.9926
16.5609 | -34.8235 | -26.0003
16.567 | -34.824 | -26.0078
16.5731 | -34.8245 | -26.0155
16.5792 | -34.825 | -26.0232
16.5853 | -34.8254 | -26.0309
16.5914 | -34.8259 | -26.0387
16.5975 | -34.8264 | -26.0465
16.6036 | -34.8269 | -26.0539
16.6097 | -34.8274 | -26.0615
16.6158 | -34.8278 | -26.0693
16.6219 | -34.8283 | -26.077
16.628 | -34.8288 | -26.0848
16.6341 | -34.8293 | -26.0925
16.6402 | -34.8297 | -26.0999
16.6463 | -34.8302 | -26.1075
16.6524 | -34.8307 | -26.1151
16.6585 | -34.8312 | -26.1229
16.6646 | -34.8317 | -26.1305
16.6707 | -34.8321 | -26.138
16.6768 | -34.8326 | -26.1458
16.6829 | -34.8331 | -26.1532
16.689 | -34.8336 | -26.1608
16.6952 | -34.834 | -26.1682
16.7013 | -34.8346 | -26.1759
16.7074 | -34.8351 | -26.1833
16.7135 | -34.8356 | -26.1908
16.7196 | -34.8361 | -26.1984
16.7257 | -34.8365 | -26.2059
16.7318 | -34.837 | -26.2135
16.7379 | -34.8375 | -26.2211
16.744 | -34.838 | -26.2285
16.7501 | -34.8385 | -26.2361
16.7562 | -34.839 | -26.2437
16.7623 | -34.8395 | -26.2513
16.7684 | -34.8399 | -26.259
16.7745 | -34.8404 | -26.2667
16.7806 | -34.8409 | -26.2742
16.7867 | -34.8415 | -26.2818
16.7928 | -34.8419 | -26.2896
16.7989 | -34.8424 | -26.2971
16.805 | -34.8429 | -26.3049
16.8111 | -34.8434 | -26.3124
16.8172 | -34.8439 | -26.3201
16.8233 | -34.8444 | -26.3277
16.8294 | -34.8449 | -26.3352
16.8356 | -34.8454 | -26.3429
16.8417 | -34.8459 | -26.3505
16.8478 | -34.8463 | -26.3581
16.8539 | -34.8468 | -26.3655
16.86 | -34.8473 | -26.3732
16.8661 | -34.8478 | -26.3807
16.8722 | -34.8483 | -26.3884
16.8783 | -34.8488 | -26.3961
16.8844 | -34.8493 | -26.4036
16.8905 | -34.8498 | -26.4113
16.8966 | -34.8503 | -26.4188
16.9027 | -34.8508 | -26.4264
16.9088 | -34.8513 | -26.4341
16.9149 | -34.8518 | -26.4419
16.921 | -34.8523 | -26.4494
16.9271 | -34.8528 | -26.457
16.9332 | -34.8532 | -26.4648
16.9393 | -34.8537 | -26.4724
16.9454 | -34.8542 | -26.4799
16.9515 | -34.8547 | -26.4876
16.9576 | -34.8552 | -26.495
16.9637 | -34.8557 | -26.5024
16.9698 | -34.8562 | -26.5099
16.9759 | -34.8567 | -26.5176
16.9821 | -34.8571 | -26.5254
16.9882 | -34.8576 | -26.5327
16.9943 | -34.8581 | -26.5402
17.0004 | -34.8586 | -26.5478
17.0065 | -34.8591 | -26.5553
17.0126 | -34.8596 | -26.5629
17.0187 | -34.8601 | -26.5707
17.0248 | -34.8606 | -26.578
17.0309 | -34.8611 | -26.5854
17.037 | -34.8615 | -26.593
17.0431 | -34.862 | -26.6005
17.0492 | -34.8625 | -26.6081
17.0553 | -34.863 | -26.6156
17.0614 | -34.8635 | -26.623
17.0675 | -34.864 | -26.6306
17.0736 | -34.8645 | -26.6383
17.0797 | -34.8649 | -26.6459
17.0858 | -34.8654 | -26.6536
17.0919 | -34.8659 | -26.661
17.098 | -34.8664 | -26.6682
17.1041 | -34.8669 | -26.6756
17.1102 | -34.8674 | -26.6832
17.1163 | -34.8679 | -26.6906
17.1225 | -34.8684 | -26.6981
17.1286 | -34.8689 | -26.7055
17.1347 | -34.8694 | -26.7131
17.1408 | -34.8698 | -26.7208
17.1469 | -34.8703 | -26.7284
17.153 | -34.8708 | -26.7358
17.1591 | -34.8713 | -26.7434
17.1652 | -34.8718 | -26.7507
17.1713 | -34.8723 | -26.7582
17.1774 | -34.8728 | -26.7658
17.1835 | -34.8732 | -26.7733
17.1896 | -34.8737 | -26.7807
17.1957 | -34.8742 | -26.7883
17.2018 | -34.8747 | -26.7956
17.2079 | -34.8752 | -26.8031
17.214 | -34.8757 | -26.8105
17.2201 | -34.8761 | -26.8181
17.2262 | -34.8766 | -26.8258
17.2323 | -34.8771 | -26.8333
17.2384 | -34.8776 | -26.8406
17.2445 | -34.878 | -26.8481
17.2506 | -34.8786 | -26.8557
17.2567 | -34.879 | -26.8633
17.2628 | -34.8795 | -26.8707
17.269 | -34.88 | -26.8781
17.2751 | -34.8805 | -26.8858
17.2812 | -34.881 | -26.8934
17.2873 | -34.8815 | -26.901
17.2934 | -34.882 | -26.9086
17.2995 | -34.8824 | -26.9161
17.3056 | -34.8829 | -26.9232
17.3117 | -34.8834 | -26.9309
17.3178 | -34.8839 | -26.9383
17.3239 | -34.8844 | -26.9457
17.33 | -34.8849 | -26.9533
17.3361 | -34.8854 | -26.9607
17.3422 | -34.8858 | -26.968
17.3483 | -34.8863 | -26.9755
17.3544 | -34.8868 | -26.983
17.3605 | -34.8873 | -26.9906
17.3666 | -34.8878 | -26.9981
17.3727 | -34.8883 | -27.0056
17.3788 | -34.8888 | -27.0128
17.3849 | -34.8892 | -27.0203
17.391 | -34.8897 | -27.0279
17.3971 | -34.8902 | -27.0354
17.4032 | -34.8907 | -27.0428
17.4094 | -34.8911 | -27.0501
17.4155 | -34.8916 | -27.0575
17.4216 | -34.8921 | -27.0648
17.4277 | -34.8926 | -27.0724
17.4338 | -34.8931 | -27.0798
17.4399 | -34.8936 | -27.0872
17.446 | -34.894 | -27.0944
17.4521 | -34.8945 | -27.1018
17.4582 | -34.895 | -27.1091
17.4643 | -34.8955 | -27.1166
17.4704 | -34.896 | -27.1241
17.4765 | -34.8964 | -27.1316
17.4826 | -34.8969 | -27.1389
17.4887 | -34.8974 | -27.1465
17.4948 | -34.8979 | -27.154
17.5009 | -34.8984 | -27.1615
17.507 | -34.8989 | -27.1689
17.5131 | -34.8994 | -27.1763
17.5192 | -34.8998 | -27.1836
17.5253 | -34.9003 | -27.191
17.5314 | -34.9007 | -27.1985
17.5375 | -34.9012 | -27.2059
17.5436 | -34.9017 | -27.2135
17.5497 | -34.9022 | -27.2209
17.5559 | -34.9026 | -27.2281
17.562 | -34.9031 | -27.2353
17.5681 | -34.9036 | -27.2427
17.5742 | -34.9041 | -27.2503
17.5803 | -34.9046 | -27.2576
17.5864 | -34.9051 | -27.2649
17.5925 | -34.9055 | -27.2724
17.5986 | -34.906 | -27.2797
17.6047 | -34.9065 | -27.2873
17.6108 | -34.907 | -27.2947
17.6169 | -34.9075 | -27.302
17.623 | -34.908 | -27.3092
17.6291 | -34.9084 | -27.3165
17.6352 | -34.9089 | -27.3238
17.6413 | -34.9094 | -27.3313
17.6474 | -34.9099 | -27.3387
17.6535 | -34.9103 | -27.3461
17.6596 | -34.9108 | -27.3534
17.6657 | -34.9113 | -27.3608
17.6718 | -34.9117 | -27.3683
17.6779 | -34.9122 | -27.3757
17.684 | -34.9127 | -27.3832
17.6901 | -34.9132 | -27.3906
17.6963 | -34.9137 | -27.3978
17.7024 | -34.9142 | -27.4051
17.7085 | -34.9146 | -27.4125
17.7146 | -34.9151 | -27.4198
17.7207 | -34.9156 | -27.4272
17.7268 | -34.9161 | -27.4342
17.7329 | -34.9165 | -27.4416
17.739 | -34.917 | -27.4489
17.7451 | -34.9175 | -27.4561
17.7512 | -34.918 | -27.4636
17.7573 | -34.9185 | -27.4711
17.7634 | -34.9189 | -27.4781
17.7695 | -34.9194 | -27.4854
17.7756 | -34.9199 | -27.4927
17.7817 | -34.9204 | -27.5001
17.7878 | -34.9209 | -27.5076
17.7939 | -34.9214 | -27.5148
17.8 | -34.9218 | -27.5221
17.8061 | -34.9223 | -27.5294
17.8122 | -34.9228 | -27.5368
17.8183 | -34.9233 | -27.5441
17.8244 | -34.9237 | -27.5514
17.8305 | -34.9242 | -27.5588
17.8366 | -34.9247 | -27.5659
17.8428 | -34.9252 | -27.5733
17.8489 | -34.9257 | -27.5806
17.855 | -34.9261 | -27.5879
17.8611 | -34.9266 | -27.5953
17.8672 | -34.9271 | -27.6023
17.8733 | -34.9276 | -27.6097
17.8794 | -34.9281 | -27.6169
17.8855 | -34.9285 | -27.6241
17.8916 | -34.929 | -27.6315
17.8977 | -34.9294 | -27.6389
17.9038 | -34.9299 | -27.6459
17.9099 | -34.9304 | -27.6532
17.916 | -34.9309 | -27.6606
17.9221 | -34.9314 | -27.6681
17.9282 | -34.9319 | -27.6753
17.9343 | -34.9323 | -27.6827
17.9404 | -34.9328 | -27.6899
17.9465 | -34.9333 | -27.6972
17.9526 | -34.9337 | -27.7045
17.9587 | -34.9342 | -27.7119
17.9648 | -34.9347 | -27.719
17.9709 | -34.9351 | -27.7264
17.977 | -34.9356 | -27.7337
17.9832 | -34.9361 | -27.7412
17.9893 | -34.9365 | -27.7486
17.9954 | -34.937 | -27.7559
18.0015 | -34.9374 | -27.7633
18.0076 | -34.9379 | -27.7704
18.0137 | -34.9384 | -27.7777
18.0198 | -34.9388 | -27.7849
18.0259 | -34.9393 | -27.7922
18.032 | -34.9397 | -27.7996
18.0381 | -34.9402 | -27.8068
18.0442 | -34.9406 | -27.814
18.0503 | -34.9411 | -27.8214
18.0564 | -34.9416 | -27.8287
18.0625 | -34.9421 | -27.8361
18.0686 | -34.9425 | -27.8435
18.0747 | -34.943 | -27.8508
18.0808 | -34.9435 | -27.858
18.0869 | -34.944 | -27.8654
18.093 | -34.9445 | -27.8726
18.0991 | -34.9449 | -27.8799
18.1052 | -34.9454 | -27.8873
18.1113 | -34.9459 | -27.8944
18.1174 | -34.9464 | -27.9016
18.1236 | -34.9468 | -27.9089
18.1297 | -34.9473 | -27.9163
18.1358 | -34.9478 | -27.9236
18.1419 | -34.9483 | -27.9308
18.148 | -34.9487 | -27.938
18.1541 | -34.9492 | -27.9449
18.1602 | -34.9497 | -27.9523
18.1663 | -34.9502 | -27.9596
18.1724 | -34.9506 | -27.9668
18.1785 | -34.9511 | -27.9741
18.1846 | -34.9515 | -27.9814
18.1907 | -34.952 | -27.9888
18.1968 | -34.9525 | -27.9961
18.2029 | -34.9529 | -28.0035
18.209 | -34.9534 | -28.0109
18.2151 | -34.9539 | -28.0183
18.2212 | -34.9543 | -28.0253
18.2273 | -34.9548 | -28.0324
18.2334 | -34.9553 | -28.0397
18.2395 | -34.9558 | -28.047
18.2456 | -34.9562 | -28.0542
18.2517 | -34.9567 | -28.0614
18.2578 | -34.9571 | -28.0683
18.2639 | -34.9576 | -28.0756
18.2701 | -34.9581 | -28.0828
18.2762 | -34.9586 | -28.0898
18.2823 | -34.9591 | -28.097
18.2884 | -34.9596 | -28.104
18.2945 | -34.96 | -28.111
18.3006 | -34.9605 | -28.1184
18.3067 | -34.961 | -28.1257
18.3128 | -34.9614 | -28.133
18.3189 | -34.9619 | -28.1404
18.325 | -34.9624 | -28.1474
18.3311 | -34.9628 | -28.1547
18.3372 | -34.9633 | -28.1619
18.3433 | -34.9638 | -28.1691
18.3494 | -34.9642 | -28.1765
18.3555 | -34.9647 | -28.1839
18.3616 | -34.9652 | -28.1908
18.3677 | -34.9657 | -28.1981
18.3738 | -34.9662 | -28.2052
18.3799 | -34.9666 | -28.2125
18.386 | -34.9671 | -28.2195
18.3921 | -34.9675 | -28.2268
18.3982 | -34.968 | -28.2338
18.4043 | -34.9684 | -28.2411
18.4105 | -34.9689 | -28.2483
18.4166 | -34.9693 | -28.2555
18.4227 | -34.9698 | -28.2629
18.4288 | -34.9702 | -28.2701
18.4349 | -34.9707 | -28.2772
18.441 | -34.9712 | -28.2845
18.4471 | -34.9716 | -28.2915
18.4532 | -34.9721 | -28.2986
18.4593 | -34.9726 | -28.3058
18.4654 | -34.973 | -28.3129
18.4715 | -34.9734 | -28.3202
18.4776 | -34.9739 | -28.3277
18.4837 | -34.9743 | -28.335
18.4898 | -34.9748 | -28.3424
18.4959 | -34.9752 | -28.3497
18.502 | -34.9757 | -28.357
18.5081 | -34.9761 | -28.3643
18.5142 | -34.9766 | -28.3715
18.5203 | -34.977 | -28.3787
18.5264 | -34.9775 | -28.3858
18.5325 | -34.9779 | -28.3929
18.5386 | -34.9784 | -28.4
18.5447 | -34.9788 | -28.4071
18.5508 | -34.9793 | -28.4144
18.557 | -34.9797 | -28.4218
18.5631 | -34.9802 | -28.429
18.5692 | -34.9806 | -28.4361
18.5753 | -34.9811 | -28.4433
18.5814 | -34.9816 | -28.4504
18.5875 | -34.982 | -28.4575
18.5936 | -34.9824 | -28.4648
18.5997 | -34.9829 | -28.472
18.6058 | -34.9833 | -28.479
18.6119 | -34.9838 | -28.4862
18.618 | -34.9842 | -28.4934
18.6241 | -34.9847 | -28.5006
18.6302 | -34.9852 | -28.5079
18.6363 | -34.9856 | -28.5152
18.6424 | -34.9861 | -28.5223
18.6485 | -34.9865 | -28.5296
18.6546 | -34.9869 | -28.5367
18.6607 | -34.9874 | -28.5438
18.6668 | -34.9879 | -28.551
18.6729 | -34.9883 | -28.5582
18.679 | -34.9888 | -28.5653
18.6851 | -34.9893 | -28.5723
18.6912 | -34.9897 | -28.5796
18.6974 | -34.9902 | -28.5867
18.7035 | -34.9906 | -28.5937
18.7096 | -34.9911 | -28.6011
18.7157 | -34.9916 | -28.608
18.7218 | -34.992 | -28.6152
18.7279 | -34.9925 | -28.6224
18.734 | -34.993 | -28.6295
18.7401 | -34.9934 | -28.6366
18.7462 | -34.9939 | -28.6434
18.7523 | -34.9944 | -28.6507
18.7584 | -34.9949 | -28.6579
18.7645 | -34.9953 | -28.6651
18.7706 | -34.9958 | -28.6721
18.7767 | -34.9962 | -28.6793
18.7828 | -34.9967 | -28.6861
18.7889 | -34.9971 | -28.6933
18.795 | -34.9976 | -28.7006
18.8011 | -34.9981 | -28.7078
18.8072 | -34.9985 | -28.7151
18.8133 | -34.999 | -28.7222
18.8194 | -34.9994 | -28.7292
18.8255 | -34.9999 | -28.7364
18.8316 | -35.0003 | -28.7435
18.8377 | -35.0008 | -28.7505
18.8439 | -35.0013 | -28.7575
18.85 | -35.0017 | -28.7646
18.8561 | -35.0022 | -28.7719
18.8622 | -35.0026 | -28.779
18.8683 | -35.0031 | -28.7862
18.8744 | -35.0036 | -28.7933
18.8805 | -35.004 | -28.8006
18.8866 | -35.0045 | -28.8079
18.8927 | -35.0049 | -28.8146
18.8988 | -35.0054 | -28.8217
18.9049 | -35.0058 | -28.8288
18.911 | -35.0063 | -28.8361
18.9171 | -35.0067 | -28.8432
18.9232 | -35.0072 | -28.8502
18.9293 | -35.0077 | -28.8573
18.9354 | -35.0081 | -28.8645
18.9415 | -35.0086 | -28.8717
18.9476 | -35.009 | -28.8787
18.9537 | -35.0095 | -28.8858
18.9598 | -35.0099 | -28.8927
18.9659 | -35.0104 | -28.8998
18.972 | -35.0109 | -28.9069
18.9781 | -35.0113 | -28.9141
18.9843 | -35.0118 | -28.921
18.9904 | -35.0122 | -28.928
18.9965 | -35.0127 | -28.9347
19.0026 | -35.0132 | -28.9418
19.0087 | -35.0136 | -28.9489
19.0148 | -35.0141 | -28.956
19.0209 | -35.0145 | -28.9633
19.027 | -35.015 | -28.9704
19.0331 | -35.0154 | -28.9773
19.0392 | -35.0159 | -28.9844
19.0453 | -35.0164 | -28.9912
19.0514 | -35.0168 | -28.9983
19.0575 | -35.0173 | -29.0056
19.0636 | -35.0177 | -29.0122
19.0697 | -35.0181 | -29.0193
19.0758 | -35.0186 | -29.0265
19.0819 | -35.019 | -29.0335
19.088 | -35.0195 | -29.0402
19.0941 | -35.0199 | -29.0473
19.1002 | -35.0203 | -29.0541
19.1063 | -35.0208 | -29.061
19.1124 | -35.0212 | -29.0682
19.1185 | -35.0217 | -29.075
19.1246 | -35.0221 | -29.0819
19.1308 | -35.0226 | -29.0892
19.1369 | -35.0231 | -29.0961
19.143 | -35.0235 | -29.1031
19.1491 | -35.0239 | -29.11
19.1552 | -35.0244 | -29.1171
19.1613 | -35.0248 | -29.1241
19.1674 | -35.0253 | -29.1313
19.1735 | -35.0257 | -29.1382
19.1796 | -35.0262 | -29.1451
19.1857 | -35.0266 | -29.1523
19.1918 | -35.0271 | -29.1595
19.1979 | -35.0275 | -29.1665
19.204 | -35.028 | -29.1732
19.2101 | -35.0284 | -29.1801
19.2162 | -35.0289 | -29.1872
19.2223 | -35.0294 | -29.1941
19.2284 | -35.0298 | -29.2013
19.2345 | -35.0303 | -29.2084
19.2406 | -35.0307 | -29.2153
19.2467 | -35.0311 | -29.2224
19.2528 | -35.0316 | -29.2295
19.2589 | -35.0321 | -29.2364
19.265 | -35.0325 | -29.2435
19.2712 | -35.033 | -29.2504
19.2773 | -35.0334 | -29.2573
19.2834 | -35.0339 | -29.2641
19.2895 | -35.0343 | -29.2711
19.2956 | -35.0347 | -29.2782
19.3017 | -35.0352 | -29.2852
19.3078 | -35.0356 | -29.2923
19.3139 | -35.0361 | -29.2992
19.32 | -35.0365 | -29.3059
19.3261 | -35.037 | -29.3131
19.3322 | -35.0374 | -29.32
19.3383 | -35.0379 | -29.327
19.3444 | -35.0383 | -29.3341
19.3505 | -35.0388 | -29.3411
19.3566 | -35.0392 | -29.3481
19.3627 | -35.0397 | -29.3551
19.3688 | -35.0401 | -29.3621
19.3749 | -35.0406 | -29.369
19.381 | -35.0411 | -29.3758
19.3871 | -35.0415 | -29.3827
19.3932 | -35.042 | -29.3894
19.3993 | -35.0424 | -29.3965
19.4054 | -35.0429 | -29.4035
19.4115 | -35.0433 | -29.4107
19.4177 | -35.0437 | -29.4176
19.4238 | -35.0442 | -29.4246
19.4299 | -35.0447 | -29.4314
19.436 | -35.0451 | -29.4385
19.4421 | -35.0456 | -29.4454
19.4482 | -35.046 | -29.4524
19.4543 | -35.0465 | -29.4592
19.4604 | -35.0469 | -29.4663
19.4665 | -35.0474 | -29.4732
19.4726 | -35.0478 | -29.4798
19.4787 | -35.0482 | -29.4869
19.4848 | -35.0487 | -29.4938
19.4909 | -35.0491 | -29.5006
19.497 | -35.0495 | -29.5074
19.5031 | -35.05 | -29.5144
19.5092 | -35.0504 | -29.5215
19.5153 | -35.0508 | -29.5283
19.5214 | -35.0513 | -29.5352
19.5275 | -35.0517 | -29.542
19.5336 | -35.0521 | -29.5488
19.5397 | -35.0526 | -29.5559
19.5458 | -35.053 | -29.563
19.5519 | -35.0535 | -29.57
19.5581 | -35.0539 | -29.5769
19.5642 | -35.0543 | -29.5841
19.5703 | -35.0547 | -29.591
19.5764 | -35.0552 | -29.5979
19.5825 | -35.0556 | -29.6047
19.5886 | -35.056 | -29.6117
19.5947 | -35.0565 | -29.6188
19.6008 | -35.0569 | -29.6258
19.6069 | -35.0574 | -29.633
19.613 | -35.0578 | -29.6401
19.6191 | -35.0583 | -29.6471
19.6252 | -35.0587 | -29.6542
19.6313 | -35.0592 | -29.6611
19.6374 | -35.0596 | -29.6683
19.6435 | -35.06 | -29.6754
19.6496 | -35.0605 | -29.6823
19.6557 | -35.0609 | -29.6894
19.6618 | -35.0614 | -29.6962
19.6679 | -35.0618 | -29.7034
19.674 | -35.0622 | -29.7104
19.6801 | -35.0627 | -29.7171
19.6862 | -35.0631 | -29.7239
19.6923 | -35.0636 | -29.7307
19.6984 | -35.064 | -29.7374
19.7046 | -35.0645 | -29.7445
19.7107 | -35.0649 | -29.7515
19.7168 | -35.0654 | -29.7585
19.7229 | -35.0658 | -29.7654
19.729 | -35.0663 | -29.7724
19.7351 | -35.0667 | -29.7788
19.7412 | -35.0671 | -29.7859
19.7473 | -35.0676 | -29.7927
19.7534 | -35.068 | -29.7999
19.7595 | -35.0684 | -29.8068
19.7656 | -35.0688 | -29.814
19.7717 | -35.0692 | -29.8209
19.7778 | -35.0697 | -29.8278
19.7839 | -35.0701 | -29.8344
19.79 | -35.0706 | -29.8411
19.7961 | -35.071 | -29.8481
19.8022 | -35.0714 | -29.8544
19.8083 | -35.0719 | -29.8614
19.8144 | -35.0724 | -29.8683
19.8205 | -35.0728 | -29.8751
19.8266 | -35.0732 | -29.8818
19.8327 | -35.0737 | -29.8886
19.8388 | -35.0741 | -29.8954
19.845 | -35.0745 | -29.9023
19.8511 | -35.0749 | -29.9092
19.8572 | -35.0754 | -29.916
19.8633 | -35.0758 | -29.9231
19.8694 | -35.0763 | -29.93
19.8755 | -35.0767 | -29.9368
19.8816 | -35.0772 | -29.9437
19.8877 | -35.0776 | -29.9504
19.8938 | -35.0781 | -29.9575
19.8999 | -35.0785 | -29.9641
19.906 | -35.0789 | -29.9711
19.9121 | -35.0794 | -29.9777
19.9182 | -35.0798 | -29.9847
19.9243 | -35.0802 | -29.9915
19.9304 | -35.0806 | -29.9983
19.9365 | -35.0811 | -30.0053
19.9426 | -35.0816 | -30.0118
19.9487 | -35.082 | -30.0187
19.9548 | -35.0825 | -30.0256
19.9609 | -35.083 | -30.0324
19.967 | -35.0834 | -30.0392
19.9731 | -35.0838 | -30.0461
19.9792 | -35.0842 | -30.0528
19.9853 | -35.0847 | -30.0597
19.9915 | -35.0851 | -30.0665
19.9976 | -35.0855 | -30.0729
20.0037 | -35.086 | -30.0797
20.0098 | -35.0864 | -30.0867
20.0159 | -35.0868 | -30.0937
20.022 | -35.0873 | -30.1006
20.0281 | -35.0878 | -30.1076
20.0342 | -35.0882 | -30.1145
20.0403 | -35.0886 | -30.1215
20.0464 | -35.0891 | -30.1282
20.0525 | -35.0895 | -30.1349
20.0586 | -35.0899 | -30.1418
20.0647 | -35.0904 | -30.1481
20.0708 | -35.0908 | -30.1548
20.0769 | -35.0913 | -30.1618
20.083 | -35.0917 | -30.1686
20.0891 | -35.0921 | -30.1755
20.0952 | -35.0926 | -30.1823
20.1013 | -35.093 | -30.1893
20.1074 | -35.0935 | -30.1961
20.1135 | -35.0939 | -30.2029
20.1196 | -35.0943 | -30.2099
20.1257 | -35.0947 | -30.2169
20.1319 | -35.0952 | -30.2236
20.138 | -35.0956 | -30.2306
20.1441 | -35.0961 | -30.2374
20.1502 | -35.0965 | -30.2441
20.1563 | -35.0969 | -30.2503
20.1624 | -35.0974 | -30.2571
20.1685 | -35.0978 | -30.2642
20.1746 | -35.0982 | -30.271
20.1807 | -35.0987 | -30.2777
20.1868 | -35.0991 | -30.2845
20.1929 | -35.0995 | -30.2908
20.199 | -35.1 | -30.2976
20.2051 | -35.1004 | -30.3043
20.2112 | -35.1009 | -30.3111
20.2173 | -35.1013 | -30.3177
20.2234 | -35.1018 | -30.3244
20.2295 | -35.1022 | -30.331
20.2356 | -35.1026 | -30.3382
20.2417 | -35.1031 | -30.3449
20.2478 | -35.1035 | -30.3515
20.2539 | -35.1039 | -30.3586
20.26 | -35.1044 | -30.3654
20.2661 | -35.1049 | -30.3721
20.2723 | -35.1053 | -30.3786
20.2784 | -35.1057 | -30.3854
20.2845 | -35.1061 | -30.3923
20.2906 | -35.1066 | -30.399
20.2967 | -35.107 | -30.4057
20.3028 | -35.1074 | -30.4125
20.3089 | -35.1079 | -30.4194
20.315 | -35.1083 | -30.4262
20.3211 | -35.1087 | -30.4331
20.3272 | -35.1092 | -30.4399
20.3333 | -35.1096 | -30.446
20.3394 | -35.11 | -30.4526
20.3455 | -35.1104 | -30.4595
20.3516 | -35.1109 | -30.4661
20.3577 | -35.1113 | -30.473
20.3638 | -35.1118 | -30.4796
20.3699 | -35.1122 | -30.4859
20.376 | -35.1126 | -30.4924
20.3821 | -35.1131 | -30.4994
20.3882 | -35.1135 | -30.5062
20.3943 | -35.1139 | -30.5128
20.4004 | -35.1143 | -30.5191
20.4065 | -35.1148 | -30.5258
20.4126 | -35.1153 | -30.5323
20.4188 | -35.1157 | -30.539
20.4249 | -35.1162 | -30.5457
20.431 | -35.1166 | -30.5521
20.4371 | -35.117 | -30.5585
20.4432 | -35.1175 | -30.5653
20.4493 | -35.1179 | -30.572
20.4554 | -35.1184 | -30.5788
20.4615 | -35.1189 | -30.5857
20.4676 | -35.1193 | -30.5923
20.4737 | -35.1197 | -30.5988
20.4798 | -35.1202 | -30.6053
20.4859 | -35.1206 | -30.6123
20.492 | -35.1211 | -30.6189
20.4981 | -35.1215 | -30.6254
20.5042 | -35.1219 | -30.632
20.5103 | -35.1223 | -30.6383
20.5164 | -35.1228 | -30.645
20.5225 | -35.1232 | -30.6517
20.5286 | -35.1236 | -30.6586
20.5347 | -35.1241 | -30.6654
20.5408 | -35.1245 | -30.672
20.5469 | -35.125 | -30.6786
20.553 | -35.1254 | -30.6851
20.5592 | -35.1258 | -30.6918
20.5653 | -35.1263 | -30.6984
20.5714 | -35.1267 | -30.7052
20.5775 | -35.1271 | -30.7116
20.5836 | -35.1276 | -30.7184
20.5897 | -35.1281 | -30.7252
20.5958 | -35.1285 | -30.7317
20.6019 | -35.1289 | -30.7385
20.608 | -35.1293 | -30.7454
20.6141 | -35.1297 | -30.7518
20.6202 | -35.1302 | -30.7583
20.6263 | -35.1307 | -30.7648
20.6324 | -35.1312 | -30.7716
20.6385 | -35.1316 | -30.7783
20.6446 | -35.1321 | -30.7853
20.6507 | -35.1325 | -30.7917
20.6568 | -35.133 | -30.7985
20.6629 | -35.1334 | -30.8053
20.669 | -35.1338 | -30.8119
20.6751 | -35.1342 | -30.8187
20.6812 | -35.1346 | -30.8249
20.6873 | -35.135 | -30.8316
20.6934 | -35.1355 | -30.8385
20.6995 | -35.1359 | -30.8452
20.7057 | -35.1363 | -30.8518
20.7118 | -35.1368 | -30.8584
20.7179 | -35.1372 | -30.865
20.724 | -35.1377 | -30.8717
20.7301 | -35.1381 | -30.8783
20.7362 | -35.1385 | -30.885
20.7423 | -35.139 | -30.8916
20.7484 | -35.1394 | -30.8982
20.7545 | -35.1399 | -30.9045
20.7606 | -35.1404 | -30.911
20.7667 | -35.1408 | -30.9172
20.7728 | -35.1412 | -30.9239
20.7789 | -35.1417 | -30.9303
20.785 | -35.1421 | -30.9369
20.7911 | -35.1426 | -30.9437
20.7972 | -35.1429 | -30.9505
20.8033 | -35.1434 | -30.9571
20.8094 | -35.1438 | -30.9637
20.8155 | -35.1442 | -30.9702
20.8216 | -35.1447 | -30.9769
20.8277 | -35.1451 | -30.9833
20.8338 | -35.1456 | -30.9896
20.8399 | -35.146 | -30.9959
20.8461 | -35.1465 | -31.0026
20.8522 | -35.147 | -31.0094
20.8583 | -35.1473 | -31.0157
20.8644 | -35.1478 | -31.0222
20.8705 | -35.1482 | -31.0289
20.8766 | -35.1487 | -31.0356
20.8827 | -35.1491 | -31.0422
20.8888 | -35.1495 | -31.049
20.8949 | -35.15 | -31.0556
20.901 | -35.1504 | -31.0623
20.9071 | -35.1508 | -31.069
20.9132 | -35.1513 | -31.0755
20.9193 | -35.1518 | -31.0818
20.9254 | -35.1522 | -31.0884
20.9315 | -35.1527 | -31.0943
20.9376 | -35.1532 | -31.1008
20.9437 | -35.1536 | -31.1077
20.9498 | -35.1541 | -31.1145
20.9559 | -35.1545 | -31.1209
20.962 | -35.1549 | -31.1275
20.9681 | -35.1554 | -31.1343
20.9742 | -35.1558 | -31.1405
20.9803 | -35.1562 | -31.147
20.9864 | -35.1566 | -31.1537
20.9926 | -35.1571 | -31.1605
20.9987 | -35.1575 | -31.1665
21.0048 | -35.1579 | -31.1729
21.0109 | -35.1584 | -31.1794
21.017 | -35.1588 | -31.1859
21.0231 | -35.1593 | -31.1922
21.0292 | -35.1597 | -31.1989
21.0353 | -35.1602 | -31.205
21.0414 | -35.1607 | -31.2117
21.0475 | -35.1611 | -31.2184
21.0536 | -35.1615 | -31.2249
21.0597 | -35.1619 | -31.2313
21.0658 | -35.1623 | -31.2379
21.0719 | -35.1627 | -31.2445
21.078 | -35.1631 | -31.2511
21.0841 | -35.1635 | -31.2579
21.0902 | -35.164 | -31.2644
21.0963 | -35.1644 | -31.271
21.1024 | -35.1649 | -31.2774
21.1085 | -35.1653 | -31.2831
21.1146 | -35.1658 | -31.2894
21.1207 | -35.1662 | -31.2962
21.1268 | -35.1666 | -31.303
21.133 | -35.1671 | -31.3097
21.1391 | -35.1674 | -31.3158
21.1452 | -35.1678 | -31.322
21.1513 | -35.1682 | -31.3286
21.1574 | -35.1687 | -31.3355
21.1635 | -35.1691 | -31.3421
21.1696 | -35.1695 | -31.3482
21.1757 | -35.1699 | -31.355
21.1818 | -35.1704 | -31.3619
21.1879 | -35.1708 | -31.3683
21.194 | -35.1713 | -31.3748
21.2001 | -35.1717 | -31.381
21.2062 | -35.1722 | -31.3874
21.2123 | -35.1726 | -31.3934
21.2184 | -35.173 | -31.4
21.2245 | -35.1735 | -31.4066
21.2306 | -35.1739 | -31.4129
21.2367 | -35.1744 | -31.4196
21.2428 | -35.1748 | -31.426
21.2489 | -35.1753 | -31.4325
21.255 | -35.1757 | -31.4393
21.2611 | -35.1762 | -31.4454
21.2672 | -35.1766 | -31.4517
21.2733 | -35.177 | -31.4581
21.2795 | -35.1774 | -31.4646
21.2856 | -35.1779 | -31.471
21.2917 | -35.1783 | -31.4778
21.2978 | -35.1787 | -31.4843
21.3039 | -35.1792 | -31.4908
21.31 | -35.1797 | -31.4977
21.3161 | -35.18 | -31.5036
21.3222 | -35.1804 | -31.5101
21.3283 | -35.1808 | -31.5165
21.3344 | -35.1813 | -31.523
21.3405 | -35.1817 | -31.5296
21.3466 | -35.1822 | -31.5362
21.3527 | -35.1827 | -31.5425
21.3588 | -35.1831 | -31.549
21.3649 | -35.1835 | -31.5555
21.371 | -35.184 | -31.5619
21.3771 | -35.1844 | -31.5682
21.3832 | -35.1849 | -31.5748
21.3893 | -35.1853 | -31.5813
21.3954 | -35.1857 | -31.5879
21.4015 | -35.1861 | -31.5943
21.4076 | -35.1866 | -31.6009
21.4137 | -35.187 | -31.6072
21.4199 | -35.1875 | -31.6135
21.426 | -35.1879 | -31.6195
21.4321 | -35.1883 | -31.6259
21.4382 | -35.1888 | -31.6325
21.4443 | -35.1892 | -31.6388
21.4504 | -35.1896 | -31.6451
21.4565 | -35.19 | -31.6514
21.4626 | -35.1905 | -31.6579
21.4687 | -35.1909 | -31.6643
21.4748 | -35.1914 | -31.6709
21.4809 | -35.1918 | -31.6772
21.487 | -35.1922 | -31.6835
21.4931 | -35.1926 | -31.6892
21.4992 | -35.1931 | -31.6957
21.5053 | -35.1936 | -31.7026
21.5114 | -35.194 | -31.7088
21.5175 | -35.1944 | -31.7151
21.5236 | -35.1948 | -31.7217
21.5297 | -35.1952 | -31.7281
21.5358 | -35.1956 | -31.7346
21.5419 | -35.1961 | -31.741
21.548 | -35.1965 | -31.7473
21.5541 | -35.1969 | -31.7537
21.5602 | -35.1973 | -31.7602
21.5664 | -35.1978 | -31.7658
21.5725 | -35.1982 | -31.7719
21.5786 | -35.1986 | -31.7785
21.5847 | -35.1991 | -31.7849
21.5908 | -35.1996 | -31.7912
21.5969 | -35.2 | -31.7973
21.603 | -35.2003 | -31.8039
21.6091 | -35.2008 | -31.8106
21.6152 | -35.2012 | -31.8169
21.6213 | -35.2017 | -31.8232
21.6274 | -35.2021 | -31.83
21.6335 | -35.2026 | -31.836
21.6396 | -35.203 | -31.8428
21.6457 | -35.2034 | -31.8495
21.6518 | -35.2038 | -31.8562
21.6579 | -35.2043 | -31.8625
21.664 | -35.2047 | -31.869
21.6701 | -35.2051 | -31.875
21.6762 | -35.2056 | -31.8813
21.6823 | -35.206 | -31.8874
21.6884 | -35.2065 | -31.894
21.6945 | -35.207 | -31.9001
21.7006 | -35.2074 | -31.9065
21.7068 | -35.2079 | -31.9129
21.7129 | -35.2083 | -31.9196
21.719 | -35.2088 | -31.9261
21.7251 | -35.2092 | -31.9321
21.7312 | -35.2096 | -31.9387
21.7373 | -35.21 | -31.945
21.7434 | -35.2104 | -31.9509
21.7495 | -35.2108 | -31.9573
21.7556 | -35.2113 | -31.9637
21.7617 | -35.2118 | -31.9702
21.7678 | -35.2122 | -31.9765
21.7739 | -35.2127 | -31.9821
21.78 | -35.2131 | -31.9886
21.7861 | -35.2136 | -31.9949
21.7922 | -35.214 | -32.0011
21.7983 | -35.2144 | -32.0074
21.8044 | -35.2149 | -32.0136
21.8105 | -35.2154 | -32.0193
21.8166 | -35.2158 | -32.0255
21.8227 | -35.2162 | -32.0317
21.8288 | -35.2167 | -32.0381
21.8349 | -35.2171 | -32.0447
21.841 | -35.2175 | -32.0508
21.8471 | -35.2179 | -32.0566
21.8533 | -35.2183 | -32.0632
21.8594 | -35.2187 | -32.0694
21.8655 | -35.2191 | -32.0758
21.8716 | -35.2195 | -32.082
21.8777 | -35.22 | -32.0879
21.8838 | -35.2205 | -32.0942
21.8899 | -35.2209 | -32.1004
21.896 | -35.2214 | -32.107
21.9021 | -35.2218 | -32.1132
21.9082 | -35.2222 | -32.1198
21.9143 | -35.2227 | -32.1261
21.9204 | -35.2231 | -32.1322
21.9265 | -35.2236 | -32.1385
21.9326 | -35.224 | -32.1449
21.9387 | -35.2245 | -32.1513
21.9448 | -35.225 | -32.1576
21.9509 | -35.2253 | -32.164
21.957 | -35.2257 | -32.1701
21.9631 | -35.2262 | -32.176
21.9692 | -35.2266 | -32.1825
21.9753 | -35.227 | -32.1891
21.9814 | -35.2275 | -32.1956
21.9875 | -35.2279 | -32.2017
21.9937 | -35.2283 | -32.208
21.9998 | -35.2288 | -32.2144
22.0059 | -35.2292 | -32.2209
22.012 | -35.2297 | -32.2272
22.0181 | -35.2301 | -32.2337
22.0242 | -35.2306 | -32.2395
22.0303 | -35.231 | -32.2456
22.0364 | -35.2314 | -32.2517
22.0425 | -35.2319 | -32.2582
22.0486 | -35.2323 | -32.2643
22.0547 | -35.2328 | -32.2702
22.0608 | -35.2332 | -32.2767
22.0669 | -35.2336 | -32.283
22.073 | -35.234 | -32.2891
22.0791 | -35.2344 | -32.2954
22.0852 | -35.2348 | -32.3016
22.0913 | -35.2352 | -32.3079
22.0974 | -35.2358 | -32.3142
22.1035 | -35.2362 | -32.3207
22.1096 | -35.2366 | -32.3266
22.1157 | -35.2371 | -32.3332
22.1218 | -35.2375 | -32.3393
22.1279 | -35.238 | -32.3452
22.134 | -35.2384 | -32.3513
22.1402 | -35.2388 | -32.3574
22.1463 | -35.2392 | -32.3639
22.1524 | -35.2397 | -32.3701
22.1585 | -35.2401 | -32.3763
22.1646 | -35.2405 | -32.382
22.1707 | -35.241 | -32.3882
22.1768 | -35.2415 | -32.3945
22.1829 | -35.2419 | -32.4007
22.189 | -35.2423 | -32.407
22.1951 | -35.2427 | -32.4131
22.2012 | -35.2432 | -32.4193
22.2073 | -35.2437 | -32.4254
22.2134 | -35.2441 | -32.4313
22.2195 | -35.2445 | -32.4376
22.2256 | -35.2449 | -32.4439
22.2317 | -35.2453 | -32.4497
22.2378 | -35.2457 | -32.4556
22.2439 | -35.2462 | -32.4621
22.25 | -35.2467 | -32.4682
22.2561 | -35.2472 | -32.474
22.2622 | -35.2477 | -32.48
22.2683 | -35.248 | -32.4857
22.2744 | -35.2484 | -32.4917
22.2806 | -35.2489 | -32.498
22.2867 | -35.2493 | -32.5042
22.2928 | -35.2498 | -32.5102
22.2989 | -35.2502 | -32.5162
22.305 | -35.2506 | -32.5223
22.3111 | -35.2511 | -32.5288
22.3172 | -35.2515 | -32.5349
22.3233 | -35.252 | -32.5412
22.3294 | -35.2524 | -32.5476
22.3355 | -35.2529 | -32.5535
22.3416 | -35.2533 | -32.5597
22.3477 | -35.2537 | -32.5662
22.3538 | -35.2542 | -32.5724
22.3599 | -35.2546 | -32.5784
22.366 | -35.255 | -32.5844
22.3721 | -35.2555 | -32.5905
22.3782 | -35.2559 | -32.5966
22.3843 | -35.2563 | -32.6028
22.3904 | -35.2568 | -32.6091
22.3965 | -35.2572 | -32.6153
22.4026 | -35.2576 | -32.6217
22.4087 | -35.2581 | -32.6277
22.4148 | -35.2585 | -32.6342
22.4209 | -35.2589 | -32.6404
22.4271 | -35.2594 | -32.6466
22.4332 | -35.2598 | -32.6527
22.4393 | -35.2603 | -32.6591
22.4454 | -35.2607 | -32.6648
22.4515 | -35.2612 | -32.671
22.4576 | -35.2615 | -32.6771
22.4637 | -35.262 | -32.683
22.4698 | -35.2624 | -32.689
22.4759 | -35.2628 | -32.6949
22.482 | -35.2633 | -32.7012
22.4881 | -35.2637 | -32.7071
22.4942 | -35.2642 | -32.7132
22.5003 | -35.2646 | -32.7194
22.5064 | -35.2651 | -32.7255
22.5125 | -35.2655 | -32.7311
22.5186 | -35.2659 | -32.7374
22.5247 | -35.2664 | -32.7438
22.5308 | -35.2668 | -32.7499
22.5369 | -35.2673 | -32.7564
22.543 | -35.2677 | -32.7628
22.5491 | -35.2682 | -32.7686
22.5552 | -35.2686 | -32.775
22.5613 | -35.269 | -32.7813
22.5675 | -35.2694 | -32.7872
22.5736 | -35.2699 | -32.7934
22.5797 | -35.2704 | -32.7999
22.5858 | -35.2709 | -32.8054
22.5919 | -35.2714 | -32.8113
22.598 | -35.2718 | -32.8173
22.6041 | -35.2722 | -32.8233
22.6102 | -35.2726 | -32.8291
22.6163 | -35.2731 | -32.8353
22.6224 | -35.2736 | -32.8416
22.6285 | -35.274 | -32.8475
22.6346 | -35.2744 | -32.8534
22.6407 | -35.2748 | -32.8597
22.6468 | -35.2752 | -32.8663
22.6529 | -35.2757 | -32.8721
22.659 | -35.276 | -32.8779
22.6651 | -35.2765 | -32.8841
22.6712 | -35.2769 | -32.8901
22.6773 | -35.2773 | -32.896
22.6834 | -35.2777 | -32.9022
22.6895 | -35.2782 | -32.9082
22.6956 | -35.2786 | -32.9144
22.7017 | -35.279 | -32.9206
22.7079 | -35.2794 | -32.9264
22.714 | -35.2799 | -32.9328
22.7201 | -35.2803 | -32.939
22.7262 | -35.2808 | -32.9448
22.7323 | -35.2812 | -32.9507
22.7384 | -35.2816 | -32.957
22.7445 | -35.2821 | -32.9633
22.7506 | -35.2825 | -32.9693
22.7567 | -35.283 | -32.9755
22.7628 | -35.2834 | -32.9811
22.7689 | -35.2838 | -32.9873
22.775 | -35.2842 | -32.9935
22.7811 | -35.2846 | -32.9996
22.7872 | -35.285 | -33.0057
22.7933 | -35.2854 | -33.0118
22.7994 | -35.2859 | -33.0178
22.8055 | -35.2864 | -33.0236
22.8116 | -35.2868 | -33.0294
22.8177 | -35.2872 | -33.0358
22.8238 | -35.2876 | -33.0415
22.8299 | -35.2879 | -33.0476
22.836 | -35.2883 | -33.0539
22.8421 | -35.2888 | -33.0602
22.8482 | -35.2892 | -33.0661
22.8544 | -35.2897 | -33.0722
22.8605 | -35.2901 | -33.0787
22.8666 | -35.2905 | -33.0835
22.8727 | -35.291 | -33.0896
22.8788 | -35.2914 | -33.0958
22.8849 | -35.2917 | -33.1017
22.891 | -35.2922 | -33.1077
22.8971 | -35.2928 | -33.1139
22.9032 | -35.2932 | -33.119
22.9093 | -35.2937 | -33.1249
22.9154 | -35.2941 | -33.1311
22.9215 | -35.2945 | -33.1374
22.9276 | -35.295 | -33.1438
22.9337 | -35.2954 | -33.1493
22.9398 | -35.2958 | -33.1553
22.9459 | -35.2962 | -33.1615
22.952 | -35.2966 | -33.1676
22.9581 | -35.2971 | -33.1737
22.9642 | -35.2975 | -33.1796
22.9703 | -35.2979 | -33.1852
22.9764 | -35.2983 | -33.1913
22.9825 | -35.2987 | -33.1974
22.9886 | -35.2992 | -33.2035
22.9948 | -35.2997 | -33.2093
23.0009 | -35.3001 | -33.2154
23.007 | -35.3005 | -33.2215
23.0131 | -35.301 | -33.2272
23.0192 | -35.3014 | -33.2334
23.0253 | -35.3018 | -33.2394
23.0314 | -35.3023 | -33.2456
23.0375 | -35.3027 | -33.2514
23.0436 | -35.3031 | -33.2569
23.0497 | -35.3035 | -33.263
23.0558 | -35.304 | -33.2691
23.0619 | -35.3044 | -33.2753
23.068 | -35.3049 | -33.2815
23.0741 | -35.3052 | -33.2873
23.0802 | -35.3056 | -33.2931
23.0863 | -35.306 | -33.2993
23.0924 | -35.3065 | -33.3053
23.0985 | -35.3069 | -33.3113
23.1046 | -35.3073 | -33.3175
23.1107 | -35.3078 | -33.3235
23.1168 | -35.3082 | -33.3296
23.1229 | -35.3086 | -33.3354
23.129 | -35.3091 | -33.3413
23.1351 | -35.3095 | -33.3472
23.1413 | -35.3099 | -33.3531
23.1474 | -35.3104 | -33.359
23.1535 | -35.3108 | -33.365
23.1596 | -35.3113 | -33.3709
23.1657 | -35.3117 | -33.3768
23.1718 | -35.3121 | -33.3825
23.1779 | -35.3125 | -33.3888
23.184 | -35.3129 | -33.3947
23.1901 | -35.3134 | -33.4007
23.1962 | -35.3138 | -33.4067
23.2023 | -35.3142 | -33.4129
23.2084 | -35.3146 | -33.419
23.2145 | -35.3151 | -33.4245
23.2206 | -35.3155 | -33.4305
23.2267 | -35.3159 | -33.4363
23.2328 | -35.3164 | -33.4423
23.2389 | -35.3168 | -33.4483
23.245 | -35.3172 | -33.4543
23.2511 | -35.3176 | -33.4602
23.2572 | -35.3181 | -33.4661
23.2633 | -35.3185 | -33.472
23.2694 | -35.3189 | -33.4783
23.2755 | -35.3193 | -33.4844
23.2817 | -35.3198 | -33.4905
23.2878 | -35.3202 | -33.4961
23.2939 | -35.3207 | -33.5023
23.3 | -35.3211 | -33.5083
23.3061 | -35.3215 | -33.5144
23.3122 | -35.3219 | -33.5203
23.3183 | -35.3224 | -33.5262
23.3244 | -35.3228 | -33.532
23.3305 | -35.3232 | -33.538
23.3366 | -35.3236 | -33.5438
23.3427 | -35.3241 | -33.5496
23.3488 | -35.3245 | -33.5555
23.3549 | -35.3249 | -33.5613
23.361 | -35.3253 | -33.5673
23.3671 | -35.3258 | -33.5729
23.3732 | -35.3262 | -33.5789
23.3793 | -35.3266 | -33.5845
23.3854 | -35.327 | -33.5904
23.3915 | -35.3275 | -33.5966
23.3976 | -35.3279 | -33.6025
23.4037 | -35.3283 | -33.6086
23.4098 | -35.3288 | -33.6145
23.4159 | -35.3292 | -33.6204
23.422 | -35.3297 | -33.6265
23.4282 | -35.3301 | -33.6325
23.4343 | -35.3305 | -33.6384
23.4404 | -35.3309 | -33.6443
23.4465 | -35.3313 | -33.6501
23.4526 | -35.3317 | -33.6559
23.4587 | -35.3321 | -33.6617
23.4648 | -35.3324 | -33.6672
23.4709 | -35.3329 | -33.6731
23.477 | -35.3333 | -33.6789
23.4831 | -35.3337 | -33.6849
23.4892 | -35.3341 | -33.6906
23.4953 | -35.3345 | -33.6966
23.5014 | -35.3349 | -33.7024
23.5075 | -35.3354 | -33.7081
23.5136 | -35.3358 | -33.7139
23.5197 | -35.3362 | -33.7198
23.5258 | -35.3367 | -33.7258
23.5319 | -35.3371 | -33.7314
23.538 | -35.3375 | -33.7373
23.5441 | -35.3379 | -33.7433
23.5502 | -35.3383 | -33.7494
23.5563 | -35.3387 | -33.7552
23.5624 | -35.3391 | -33.7611
23.5686 | -35.3395 | -33.7666
23.5747 | -35.3399 | -33.7727
23.5808 | -35.3403 | -33.7787
23.5869 | -35.3407 | -33.7844
23.593 | -35.3412 | -33.7904
23.5991 | -35.3416 | -33.7964
23.6052 | -35.342 | -33.8018
23.6113 | -35.3424 | -33.8078
23.6174 | -35.3428 | -33.8139
23.6235 | -35.3433 | -33.8197
23.6296 | -35.3437 | -33.8258
23.6357 | -35.3441 | -33.8322
23.6418 | -35.3446 | -33.8381
23.6479 | -35.345 | -33.8439
23.654 | -35.3455 | -33.8498
23.6601 | -35.3458 | -33.8555
23.6662 | -35.3462 | -33.8615
23.6723 | -35.3467 | -33.8672
23.6784 | -35.347 | -33.8731
23.6845 | -35.3475 | -33.8791
23.6906 | -35.3479 | -33.8845
23.6967 | -35.3484 | -33.8903
23.7028 | -35.3488 | -33.8963
23.7089 | -35.3492 | -33.9017
23.7151 | -35.3496 | -33.9078
23.7212 | -35.3501 | -33.914
23.7273 | -35.3505 | -33.9199
23.7334 | -35.3508 | -33.9258
23.7395 | -35.3513 | -33.9318
23.7456 | -35.3517 | -33.9373
23.7517 | -35.3521 | -33.9433
23.7578 | -35.3525 | -33.9492
23.7639 | -35.3529 | -33.9549
23.77 | -35.3533 | -33.9607
23.7761 | -35.3538 | -33.9667
23.7822 | -35.3541 | -33.972
23.7883 | -35.3545 | -33.9779
23.7944 | -35.355 | -33.9838
23.8005 | -35.3554 | -33.9897
23.8066 | -35.3558 | -33.9957
23.8127 | -35.3562 | -34.0016
23.8188 | -35.3567 | -34.0078
23.8249 | -35.3571 | -34.0134
23.831 | -35.3574 | -34.0191
23.8371 | -35.3579 | -34.0249
23.8432 | -35.3583 | -34.0305
23.8493 | -35.3587 | -34.0362
23.8555 | -35.3591 | -34.0423
23.8616 | -35.3595 | -34.0478
23.8677 | -35.36 | -34.0536
23.8738 | -35.3604 | -34.0593
23.8799 | -35.3608 | -34.0652
23.886 | -35.3613 | -34.0705
23.8921 | -35.3617 | -34.0765
23.8982 | -35.3621 | -34.0826
23.9043 | -35.3625 | -34.0884
23.9104 | -35.3629 | -34.0941
23.9165 | -35.3634 | -34.0998
23.9226 | -35.3638 | -34.1051
23.9287 | -35.3642 | -34.1111
23.9348 | -35.3646 | -34.117
23.9409 | -35.365 | -34.1228
23.947 | -35.3654 | -34.1286
23.9531 | -35.3658 | -34.1343
23.9592 | -35.3662 | -34.1407
23.9653 | -35.3666 | -34.1462
23.9714 | -35.367 | -34.1519
23.9775 | -35.3674 | -34.1578
23.9836 | -35.3678 | -34.1638
23.9897 | -35.3681 | -34.1695
23.9958 | -35.3685 | -34.1753
24.002 | -35.369 | -34.1813
24.0081 | -35.3694 | -34.1869
24.0142 | -35.3698 | -34.1924
24.0203 | -35.3702 | -34.1985
24.0264 | -35.3705 | -34.2044
24.0325 | -35.3709 | -34.2102
24.0386 | -35.3713 | -34.2161
24.0447 | -35.3717 | -34.2221
24.0508 | -35.3721 | -34.2281
24.0569 | -35.3725 | -34.2339
24.063 | -35.3729 | -34.2399
24.0691 | -35.3733 | -34.2458
24.0752 | -35.3737 | -34.2515
24.0813 | -35.3742 | -34.2573
24.0874 | -35.3746 | -34.2631
24.0935 | -35.3749 | -34.2688
24.0996 | -35.3753 | -34.274
24.1057 | -35.3757 | -34.2798
24.1118 | -35.3761 | -34.2858
24.1179 | -35.3765 | -34.2916
24.124 | -35.3769 | -34.2975
24.1301 | -35.3773 | -34.3029
24.1362 | -35.3777 | -34.3088
24.1424 | -35.3781 | -34.3145
24.1485 | -35.3785 | -34.3205
24.1546 | -35.3789 | -34.3262
24.1607 | -35.3793 | -34.332
24.1668 | -35.3796 | -34.3376
24.1729 | -35.3801 | -34.3432
24.179 | -35.3805 | -34.349
24.1851 | -35.3809 | -34.3548
24.1912 | -35.3813 | -34.3604
24.1973 | -35.3817 | -34.3667
24.2034 | -35.3821 | -34.372
24.2095 | -35.3825 | -34.3777
24.2156 | -35.383 | -34.3836
24.2217 | -35.3834 | -34.3893
24.2278 | -35.3838 | -34.3951
24.2339 | -35.3843 | -34.4009
24.24 | -35.3846 | -34.406
24.2461 | -35.385 | -34.4118
24.2522 | -35.3855 | -34.4179
24.2583 | -35.3859 | -34.4238
24.2644 | -35.3864 | -34.4299
24.2705 | -35.3868 | -34.4357
24.2766 | -35.3872 | -34.4417
24.2827 | -35.3876 | -34.4475
24.2889 | -35.3881 | -34.4531
24.295 | -35.3884 | -34.459
24.3011 | -35.3888 | -34.4648
24.3072 | -35.3892 | -34.4706
24.3133 | -35.3896 | -34.4763
24.3194 | -35.3899 | -34.482
24.3255 | -35.3903 | -34.4876
24.3316 | -35.3908 | -34.4934
24.3377 | -35.3911 | -34.4993
24.3438 | -35.3916 | -34.5049
24.3499 | -35.3919 | -34.5109
24.356 | -35.3923 | -34.5166
24.3621 | -35.3927 | -34.5221
24.3682 | -35.393 | -34.5282
24.3743 | -35.3934 | -34.5337
24.3804 | -35.3938 | -34.5396
24.3865 | -35.3941 | -34.545
24.3926 | -35.3945 | -34.5511
24.3987 | -35.3949 | -34.5567
24.4048 | -35.3953 | -34.5621
24.4109 | -35.3957 | -34.5677
24.417 | -35.3961 | -34.5735
24.4231 | -35.3965 | -34.5791
24.4293 | -35.3968 | -34.5844
24.4354 | -35.3972 | -34.5901
24.4415 | -35.3976 | -34.5955
24.4476 | -35.3979 | -34.6012
24.4537 | -35.3984 | -34.6069
24.4598 | -35.3988 | -34.6128
24.4659 | -35.3991 | -34.6185
24.472 | -35.3995 | -34.6239
24.4781 | -35.3998 | -34.6296
24.4842 | -35.4003 | -34.6353
24.4903 | -35.4008 | -34.6413
24.4964 | -35.4012 | -34.6467
24.5025 | -35.4016 | -34.6521
24.5086 | -35.402 | -34.6581
24.5147 | -35.4023 | -34.6636
24.5208 | -35.4027 | -34.6697
24.5269 | -35.4031 | -34.6757
24.533 | -35.4036 | -34.6812
24.5391 | -35.404 | -34.6869
24.5452 | -35.4044 | -34.6929
24.5513 | -35.4048 | -34.6992
24.5574 | -35.4051 | -34.7047
24.5635 | -35.4056 | -34.7103
24.5696 | -35.406 | -34.7162
24.5758 | -35.4064 | -34.7218
24.5819 | -35.4067 | -34.7275
24.588 | -35.407 | -34.7328
24.5941 | -35.4074 | -34.7387
24.6002 | -35.4079 | -34.7444
24.6063 | -35.4083 | -34.7501
24.6124 | -35.4086 | -34.7559
24.6185 | -35.409 | -34.7615
24.6246 | -35.4094 | -34.7665
24.6307 | -35.4097 | -34.7721
24.6368 | -35.4101 | -34.7782
24.6429 | -35.4105 | -34.784
24.649 | -35.4109 | -34.7897
24.6551 | -35.4113 | -34.795
24.6612 | -35.4117 | -34.8003
24.6673 | -35.4121 | -34.8059
24.6734 | -35.4124 | -34.8115
24.6795 | -35.4128 | -34.817
24.6856 | -35.4132 | -34.8229
24.6917 | -35.4136 | -34.8283
24.6978 | -35.414 | -34.8342
24.7039 | -35.4144 | -34.8396
24.71 | -35.4148 | -34.8453
24.7162 | -35.4152 | -34.8511
24.7223 | -35.4155 | -34.8567
24.7284 | -35.4159 | -34.8624
24.7345 | -35.4163 | -34.8682
24.7406 | -35.4167 | -34.874
24.7467 | -35.4171 | -34.8796
24.7528 | -35.4175 | -34.8852
24.7589 | -35.4179 | -34.8907
24.765 | -35.4183 | -34.8961
24.7711 | -35.4187 | -34.9017
24.7772 | -35.4191 | -34.907
24.7833 | -35.4195 | -34.9127
24.7894 | -35.4198 | -34.9183
24.7955 | -35.4202 | -34.9243
24.8016 | -35.4205 | -34.9297
24.8077 | -35.4209 | -34.9351
24.8138 | -35.4213 | -34.9408
24.8199 | -35.4216 | -34.9461
24.826 | -35.4221 | -34.9516
24.8321 | -35.4225 | -34.957
24.8382 | -35.4228 | -34.9624
24.8443 | -35.4232 | -34.9681
24.8504 | -35.4236 | -34.9733
24.8565 | -35.424 | -34.979
24.8627 | -35.4244 | -34.985
24.8688 | -35.4247 | -34.9905
24.8749 | -35.425 | -34.9963
24.881 | -35.4254 | -35.0018
24.8871 | -35.4259 | -35.0076
24.8932 | -35.4263 | -35.0135
24.8993 | -35.4267 | -35.0192
24.9054 | -35.4271 | -35.0248
24.9115 | -35.4275 | -35.0307
24.9176 | -35.4279 | -35.0362
24.9237 | -35.4283 | -35.0418
24.9298 | -35.4286 | -35.0472
24.9359 | -35.429 | -35.053
24.942 | -35.4294 | -35.0585
24.9481 | -35.4298 | -35.0642
24.9542 | -35.4301 | -35.0702
24.9603 | -35.4305 | -35.0759
24.9664 | -35.4308 | -35.0816
24.9725 | -35.4312 | -35.0873
24.9786 | -35.4316 | -35.0927
24.9847 | -35.432 | -35.0983
24.9908 | -35.4323 | -35.1041
24.9969 | -35.4327 | -35.1099
25.0031 | -35.433 | -35.1157
25.0092 | -35.4333 | -35.1212
25.0153 | -35.4337 | -35.1267
25.0214 | -35.4341 | -35.1322
25.0275 | -35.4345 | -35.1379
25.0336 | -35.4349 | -35.1435
25.0397 | -35.4353 | -35.1491
25.0458 | -35.4356 | -35.1547
25.0519 | -35.436 | -35.1605
25.058 | -35.4364 | -35.1666
25.0641 | -35.4368 | -35.1723
25.0702 | -35.4371 | -35.1778
25.0763 | -35.4375 | -35.1834
25.0824 | -35.4378 | -35.1891
25.0885 | -35.4382 | -35.1948
25.0946 | -35.4385 | -35.2004
25.1007 | -35.4389 | -35.2065
25.1068 | -35.4392 | -35.212
25.1129 | -35.4397 | -35.217
25.119 | -35.4399 | -35.2225
25.1251 | -35.4403 | -35.2279
25.1312 | -35.4407 | -35.2337
25.1373 | -35.4411 | -35.2396
25.1435 | -35.4414 | -35.2451
25.1496 | -35.4418 | -35.2505
25.1557 | -35.4422 | -35.2562
25.1618 | -35.4426 | -35.2616
25.1679 | -35.443 | -35.2673
25.174 | -35.4433 | -35.2729
25.1801 | -35.4437 | -35.2785
25.1862 | -35.444 | -35.2842
25.1923 | -35.4443 | -35.29
25.1984 | -35.4446 | -35.2953
25.2045 | -35.445 | -35.3007
25.2106 | -35.4453 | -35.3068
25.2167 | -35.4457 | -35.3123
25.2228 | -35.446 | -35.3172
25.2289 | -35.4463 | -35.3225
25.235 | -35.4467 | -35.3277
25.2411 | -35.4471 | -35.3331
25.2472 | -35.4474 | -35.3392
25.2533 | -35.4477 | -35.3446
25.2594 | -35.4481 | -35.35
25.2655 | -35.4484 | -35.3557
25.2716 | -35.4488 | -35.3613
25.2777 | -35.4491 | -35.3671
25.2838 | -35.4496 | -35.3725
25.29 | -35.45 | -35.3777
25.2961 | -35.4504 | -35.3827
25.3022 | -35.4507 | -35.3881
25.3083 | -35.4511 | -35.3939
25.3144 | -35.4514 | -35.3996
25.3205 | -35.4518 | -35.4051
25.3266 | -35.4521 | -35.4102
25.3327 | -35.4524 | -35.4158
25.3388 | -35.4528 | -35.4215
25.3449 | -35.4532 | -35.4272
25.351 | -35.4535 | -35.4328
25.3571 | -35.4539 | -35.4381
25.3632 | -35.4542 | -35.4435
25.3693 | -35.4546 | -35.4493
25.3754 | -35.4549 | -35.4552
25.3815 | -35.4552 | -35.4611
25.3876 | -35.4555 | -35.4667
25.3937 | -35.4558 | -35.4724
25.3998 | -35.4563 | -35.4783
25.4059 | -35.4566 | -35.4841
25.412 | -35.457 | -35.4898
25.4181 | -35.4574 | -35.4952
25.4242 | -35.4578 | -35.5008
25.4304 | -35.4581 | -35.506
25.4365 | -35.4584 | -35.5116
25.4426 | -35.4588 | -35.5174
25.4487 | -35.4591 | -35.5229
25.4548 | -35.4594 | -35.5288
25.4609 | -35.4598 | -35.5343
25.467 | -35.4603 | -35.5394
25.4731 | -35.4606 | -35.5449
25.4792 | -35.4609 | -35.5509
25.4853 | -35.4613 | -35.5566
25.4914 | -35.4616 | -35.5626
25.4975 | -35.462 | -35.5682
25.5036 | -35.4623 | -35.5737
25.5097 | -35.4627 | -35.5794
25.5158 | -35.4631 | -35.585
25.5219 | -35.4634 | -35.5909
25.528 | -35.4637 | -35.596
25.5341 | -35.4641 | -35.6017
25.5402 | -35.4645 | -35.6071
25.5463 | -35.4649 | -35.6125
25.5524 | -35.4652 | -35.6181
25.5585 | -35.4656 | -35.6235
25.5646 | -35.466 | -35.6289
25.5707 | -35.4663 | -35.6344
25.5769 | -35.4666 | -35.6403
25.583 | -35.4669 | -35.6455
25.5891 | -35.4673 | -35.6509
25.5952 | -35.4676 | -35.6564
25.6013 | -35.468 | -35.6619
25.6074 | -35.4683 | -35.6671
25.6135 | -35.4687 | -35.6724
25.6196 | -35.469 | -35.6778
25.6257 | -35.4694 | -35.6833
25.6318 | -35.4697 | -35.6891
25.6379 | -35.4701 | -35.6946
25.644 | -35.4704 | -35.7
25.6501 | -35.4708 | -35.7056
25.6562 | -35.4711 | -35.7113
25.6623 | -35.4715 | -35.7165
25.6684 | -35.4718 | -35.7221
25.6745 | -35.4721 | -35.7275
25.6806 | -35.4725 | -35.7324
25.6867 | -35.4729 | -35.7378
25.6928 | -35.4732 | -35.7431
25.6989 | -35.4735 | -35.7488
25.705 | -35.4738 | -35.7545
25.7111 | -35.4742 | -35.7599
25.7173 | -35.4745 | -35.7651
25.7234 | -35.4749 | -35.7707
25.7295 | -35.4752 | -35.7759
25.7356 | -35.4755 | -35.7814
25.7417 | -35.4759 | -35.7867
25.7478 | -35.4762 | -35.7918
25.7539 | -35.4765 | -35.7971
25.76 | -35.4768 | -35.8024
25.7661 | -35.4772 | -35.8073
25.7722 | -35.4775 | -35.8125
25.7783 | -35.4779 | -35.8183
25.7844 | -35.4782 | -35.8235
25.7905 | -35.4785 | -35.829
25.7966 | -35.4788 | -35.835
25.8027 | -35.4791 | -35.8406
25.8088 | -35.4795 | -35.8462
25.8149 | -35.4798 | -35.8519
25.821 | -35.4801 | -35.8573
25.8271 | -35.4804 | -35.8627
25.8332 | -35.4807 | -35.8682
25.8393 | -35.481 | -35.8738
25.8454 | -35.4814 | -35.8794
25.8515 | -35.4817 | -35.8847
25.8576 | -35.482 | -35.8905
25.8638 | -35.4823 | -35.8958
25.8699 | -35.4827 | -35.9012
25.876 | -35.483 | -35.9065
25.8821 | -35.4833 | -35.912
25.8882 | -35.4836 | -35.9175
25.8943 | -35.4839 | -35.9227
25.9004 | -35.4843 | -35.9281
25.9065 | -35.4847 | -35.9334
25.9126 | -35.485 | -35.939
25.9187 | -35.4854 | -35.9441
25.9248 | -35.4857 | -35.9493
25.9309 | -35.4861 | -35.9544
25.937 | -35.4864 | -35.9603
25.9431 | -35.4868 | -35.9657
25.9492 | -35.4871 | -35.9711
25.9553 | -35.4874 | -35.9767
25.9614 | -35.4876 | -35.982
25.9675 | -35.488 | -35.9879
25.9736 | -35.4884 | -35.9934
25.9797 | -35.4887 | -35.9991
25.9858 | -35.489 | -36.0045
25.9919 | -35.4893 | -36.0101
25.998 | -35.4896 | -36.0154
26.0042 | -35.4899 | -36.0209
26.0103 | -35.4903 | -36.0262
26.0164 | -35.4906 | -36.0317
26.0225 | -35.4909 | -36.0371
26.0286 | -35.4912 | -36.043
26.0347 | -35.4915 | -36.0483
26.0408 | -35.4918 | -36.0538
26.0469 | -35.4921 | -36.0589
26.053 | -35.4924 | -36.0643
26.0591 | -35.4927 | -36.0695
26.0652 | -35.493 | -36.0745
26.0713 | -35.4933 | -36.0801
26.0774 | -35.4937 | -36.0856
26.0835 | -35.494 | -36.0912
26.0896 | -35.4944 | -36.0964
26.0957 | -35.4947 | -36.1013
26.1018 | -35.495 | -36.1063
26.1079 | -35.4953 | -36.1119
26.114 | -35.4956 | -36.117
26.1201 | -35.4959 | -36.1228
26.1262 | -35.4963 | -36.128
26.1323 | -35.4965 | -36.1335
26.1384 | -35.4969 | -36.1385
26.1445 | -35.4972 | -36.144
26.1507 | -35.4974 | -36.1493
26.1568 | -35.4977 | -36.1548
26.1629 | -35.498 | -36.1599
26.169 | -35.4984 | -36.1654
26.1751 | -35.4987 | -36.1702
26.1812 | -35.4991 | -36.1757
26.1873 | -35.4994 | -36.1811
26.1934 | -35.4998 | -36.1864
26.1995 | -35.5001 | -36.1917
26.2056 | -35.5004 | -36.1967
26.2117 | -35.5007 | -36.2021
26.2178 | -35.5009 | -36.2074
26.2239 | -35.5012 | -36.2124
26.23 | -35.5015 | -36.2178
26.2361 | -35.5018 | -36.2231
26.2422 | -35.5021 | -36.2286
26.2483 | -35.5024 | -36.234
26.2544 | -35.5027 | -36.2391
26.2605 | -35.503 | -36.2447
26.2666 | -35.5033 | -36.25
26.2727 | -35.5036 | -36.2553
26.2788 | -35.5039 | -36.2606
26.2849 | -35.5041 | -36.2657
26.2911 | -35.5045 | -36.2713
26.2972 | -35.5048 | -36.2764
26.3033 | -35.5052 | -36.2821
26.3094 | -35.5055 | -36.2875
26.3155 | -35.5058 | -36.2933
26.3216 | -35.5061 | -36.2991
26.3277 | -35.5064 | -36.3043
26.3338 | -35.5067 | -36.3096
26.3399 | -35.507 | -36.3148
26.346 | -35.5073 | -36.3198
26.3521 | -35.5076 | -36.3251
26.3582 | -35.508 | -36.3305
26.3643 | -35.5083 | -36.3358
26.3704 | -35.5087 | -36.341
26.3765 | -35.509 | -36.3463
26.3826 | -35.5093 | -36.3515
26.3887 | -35.5096 | -36.3565
26.3948 | -35.5099 | -36.3618
26.4009 | -35.5103 | -36.367
26.407 | -35.5106 | -36.3725
26.4131 | -35.5109 | -36.3777
26.4192 | -35.5112 | -36.3825
26.4253 | -35.5114 | -36.3881
26.4314 | -35.5117 | -36.3934
26.4376 | -35.5121 | -36.3988
26.4437 | -35.5124 | -36.4044
26.4498 | -35.5127 | -36.4097
26.4559 | -35.513 | -36.4154
26.462 | -35.5133 | -36.4209
26.4681 | -35.5136 | -36.426
26.4742 | -35.5139 | -36.4315
26.4803 | -35.5142 | -36.4366
26.4864 | -35.5145 | -36.4418
26.4925 | -35.5149 | -36.447
26.4986 | -35.5151 | -36.4524
26.5047 | -35.5154 | -36.4576
26.5108 | -35.5157 | -36.4629
26.5169 | -35.516 | -36.4682
26.523 | -35.5163 | -36.4735
26.5291 | -35.5166 | -36.4789
26.5352 | -35.5168 | -36.484
26.5413 | -35.5171 | -36.4889
26.5474 | -35.5174 | -36.4945
26.5535 | -35.5176 | -36.5
26.5596 | -35.518 | -36.505
26.5657 | -35.5183 | -36.5103
26.5718 | -35.5187 | -36.5153
26.578 | -35.5191 | -36.52
26.5841 | -35.5194 | -36.5253
26.5902 | -35.5198 | -36.5306
26.5963 | -35.52 | -36.5357
26.6024 | -35.5203 | -36.5409
26.6085 | -35.5206 | -36.5461
26.6146 | -35.5209 | -36.5518
26.6207 | -35.5212 | -36.5574
26.6268 | -35.5214 | -36.5631
26.6329 | -35.5218 | -36.568
26.639 | -35.5221 | -36.5735
26.6451 | -35.5225 | -36.5788
26.6512 | -35.5228 | -36.5841
26.6573 | -35.523 | -36.589
26.6634 | -35.5233 | -36.5945
26.6695 | -35.5236 | -36.5994
26.6756 | -35.5239 | -36.6049
26.6817 | -35.5242 | -36.61
26.6878 | -35.5246 | -36.6151
26.6939 | -35.5249 | -36.6205
26.7 | -35.5252 | -36.6254
26.7061 | -35.5255 | -36.6306
26.7122 | -35.5258 | -36.6357
26.7183 | -35.526 | -36.6411
26.7245 | -35.5263 | -36.6462
26.7306 | -35.5267 | -36.6512
26.7367 | -35.5269 | -36.6567
26.7428 | -35.5272 | -36.662
26.7489 | -35.5275 | -36.6671
26.755 | -35.5278 | -36.6729
26.7611 | -35.5281 | -36.6783
26.7672 | -35.5285 | -36.6837
26.7733 | -35.5287 | -36.6893
26.7794 | -35.529 | -36.6949
26.7855 | -35.5292 | -36.6996
26.7916 | -35.5295 | -36.7042
26.7977 | -35.5299 | -36.7094
26.8038 | -35.5301 | -36.7149
26.8099 | -35.5304 | -36.7202
26.816 | -35.5307 | -36.7255
26.8221 | -35.5311 | -36.7303
26.8282 | -35.5313 | -36.7351
26.8343 | -35.5316 | -36.7402
26.8404 | -35.5318 | -36.7452
26.8465 | -35.5321 | -36.7509
26.8526 | -35.5323 | -36.7562
26.8587 | -35.5327 | -36.7611
26.8649 | -35.5329 | -36.7661
26.871 | -35.5332 | -36.7713
26.8771 | -35.5336 | -36.7765
26.8832 | -35.5338 | -36.7817
26.8893 | -35.534 | -36.7871
26.8954 | -35.5343 | -36.7926
26.9015 | -35.5346 | -36.7978
26.9076 | -35.5349 | -36.8029
26.9137 | -35.5352 | -36.8073
26.9198 | -35.5355 | -36.8127
26.9259 | -35.5358 | -36.8179
26.932 | -35.536 | -36.8231
26.9381 | -35.5363 | -36.8283
26.9442 | -35.5365 | -36.8337
26.9503 | -35.5368 | -36.8393
26.9564 | -35.5371 | -36.8447
26.9625 | -35.5375 | -36.8501
26.9686 | -35.5378 | -36.8559
26.9747 | -35.5381 | -36.861
26.9808 | -35.5384 | -36.8656
26.9869 | -35.5386 | -36.8709
26.993 | -35.539 | -36.8757
26.9991 | -35.5393 | -36.8813
27.0052 | -35.5395 | -36.8866
27.0114 | -35.5398 | -36.8915
27.0175 | -35.54 | -36.8957
27.0236 | -35.5403 | -36.901
27.0297 | -35.5406 | -36.9063
27.0358 | -35.5409 | -36.912
27.0419 | -35.5411 | -36.9171
27.048 | -35.5414 | -36.9222
27.0541 | -35.5416 | -36.9272
27.0602 | -35.5418 | -36.9326
27.0663 | -35.542 | -36.9382
27.0724 | -35.5423 | -36.9431
27.0785 | -35.5426 | -36.9483
27.0846 | -35.5429 | -36.9531
27.0907 | -35.5432 | -36.958
27.0968 | -35.5435 | -36.963
27.1029 | -35.5437 | -36.9682
27.109 | -35.5441 | -36.9731
27.1151 | -35.5444 | -36.9785
27.1212 | -35.5446 | -36.9826
27.1273 | -35.545 | -36.9878
27.1334 | -35.5453 | -36.9925
27.1395 | -35.5456 | -36.9978
27.1456 | -35.5459 | -37.0027
27.1518 | -35.5461 | -37.0075
27.1579 | -35.5464 | -37.0122
27.164 | -35.5467 | -37.0175
27.1701 | -35.5471 | -37.0226
27.1762 | -35.5473 | -37.0277
27.1823 | -35.5476 | -37.0329
27.1884 | -35.5479 | -37.038
27.1945 | -35.5481 | -37.0427
27.2006 | -35.5485 | -37.0477
27.2067 | -35.5487 | -37.0529
27.2128 | -35.549 | -37.0584
27.2189 | -35.5492 | -37.0636
27.225 | -35.5495 | -37.069
27.2311 | -35.5497 | -37.0739
27.2372 | -35.55 | -37.0789
27.2433 | -35.5503 | -37.0839
27.2494 | -35.5506 | -37.0891
27.2555 | -35.5508 | -37.094
27.2616 | -35.5511 | -37.0987
27.2677 | -35.5514 | -37.1037
27.2738 | -35.5516 | -37.1094
27.2799 | -35.552 | -37.1149
27.286 | -35.5523 | -37.1203
27.2921 | -35.5525 | -37.1259
27.2983 | -35.5528 | -37.1307
27.3044 | -35.5531 | -37.1357
27.3105 | -35.5533 | -37.1404
27.3166 | -35.5536 | -37.1459
27.3227 | -35.5538 | -37.1509
27.3288 | -35.5542 | -37.1561
27.3349 | -35.5544 | -37.1608
27.341 | -35.5547 | -37.166
27.3471 | -35.555 | -37.1708
27.3532 | -35.5553 | -37.1764
27.3593 | -35.5556 | -37.1816
27.3654 | -35.5559 | -37.1866
27.3715 | -35.5561 | -37.1916
27.3776 | -35.5564 | -37.1974
27.3837 | -35.5567 | -37.2027
27.3898 | -35.557 | -37.2081
27.3959 | -35.5572 | -37.2128
27.402 | -35.5574 | -37.217
27.4081 | -35.5577 | -37.2218
27.4142 | -35.558 | -37.2262
27.4203 | -35.5583 | -37.231
27.4264 | -35.5586 | -37.2363
27.4325 | -35.5589 | -37.2415
27.4387 | -35.5592 | -37.2467
27.4448 | -35.5595 | -37.2516
27.4509 | -35.5598 | -37.257
27.457 | -35.5601 | -37.2623
27.4631 | -35.5604 | -37.2676
27.4692 | -35.5606 | -37.2724
27.4753 | -35.5608 | -37.2771
27.4814 | -35.561 | -37.2825
27.4875 | -35.5612 | -37.2878
27.4936 | -35.5615 | -37.2928
27.4997 | -35.5619 | -37.2978
27.5058 | -35.5621 | -37.3033
27.5119 | -35.5623 | -37.3086
27.518 | -35.5626 | -37.3133
27.5241 | -35.5628 | -37.3181
27.5302 | -35.563 | -37.3235
27.5363 | -35.5632 | -37.3285
27.5424 | -35.5636 | -37.3327
27.5485 | -35.5639 | -37.3373
27.5546 | -35.5641 | -37.3421
27.5607 | -35.5645 | -37.3469
27.5668 | -35.5647 | -37.352
27.5729 | -35.5651 | -37.3577
27.5791 | -35.5653 | -37.3632
27.5852 | -35.5656 | -37.3685
27.5913 | -35.5659 | -37.3735
27.5974 | -35.5662 | -37.3783
27.6035 | -35.5664 | -37.3829
27.6096 | -35.5667 | -37.3879
27.6157 | -35.5671 | -37.3918
27.6218 | -35.5674 | -37.3972
27.6279 | -35.5676 | -37.4024
27.634 | -35.5679 | -37.4072
27.6401 | -35.5681 | -37.4121
27.6462 | -35.5683 | -37.4181
27.6523 | -35.5686 | -37.4236
27.6584 | -35.5689 | -37.4285
27.6645 | -35.5691 | -37.4335
27.6706 | -35.5694 | -37.4391
27.6767 | -35.5697 | -37.4445
27.6828 | -35.5699 | -37.4495
27.6889 | -35.5702 | -37.454
27.695 | -35.5704 | -37.4593
27.7011 | -35.5706 | -37.4643
27.7072 | -35.5709 | -37.4693
27.7133 | -35.5712 | -37.4746
27.7194 | -35.5714 | -37.4794
27.7256 | -35.5717 | -37.4842
27.7317 | -35.5719 | -37.4889
27.7378 | -35.5722 | -37.4936
27.7439 | -35.5724 | -37.4988
27.75 | -35.5726 | -37.5041
27.7561 | -35.5728 | -37.5087
27.7622 | -35.5731 | -37.5137
27.7683 | -35.5734 | -37.5183
27.7744 | -35.5737 | -37.5235
27.7805 | -35.5739 | -37.5283
27.7866 | -35.5742 | -37.533
27.7927 | -35.5743 | -37.5383
27.7988 | -35.5747 | -37.5432
27.8049 | -35.575 | -37.5479
27.811 | -35.5752 | -37.5532
27.8171 | -35.5756 | -37.5581
27.8232 | -35.5759 | -37.5621
27.8293 | -35.5761 | -37.5675
27.8354 | -35.5764 | -37.5725
27.8415 | -35.5767 | -37.577
27.8476 | -35.577 | -37.5816
27.8537 | -35.5772 | -37.5867
27.8598 | -35.5774 | -37.5914
27.866 | -35.5777 | -37.5966
27.8721 | -35.578 | -37.6018
27.8782 | -35.5783 | -37.6064
27.8843 | -35.5785 | -37.6119
27.8904 | -35.5788 | -37.617
27.8965 | -35.5791 | -37.6216
27.9026 | -35.5794 | -37.6261
27.9087 | -35.5797 | -37.6307
27.9148 | -35.5799 | -37.6355
27.9209 | -35.5802 | -37.6406
27.927 | -35.5805 | -37.6451
27.9331 | -35.5807 | -37.6497
27.9392 | -35.581 | -37.6553
27.9453 | -35.5813 | -37.6606
27.9514 | -35.5815 | -37.6662
27.9575 | -35.5818 | -37.6711
27.9636 | -35.582 | -37.6767
27.9697 | -35.5823 | -37.6817
27.9758 | -35.5824 | -37.6862
27.9819 | -35.5828 | -37.6909
27.988 | -35.583 | -37.6959
27.9941 | -35.5832 | -37.7009
28.0002 | -35.5833 | -37.7065
28.0063 | -35.5835 | -37.7116
28.0125 | -35.5837 | -37.7166
28.0186 | -35.584 | -37.7222
28.0247 | -35.5843 | -37.7274
28.0308 | -35.5846 | -37.732
28.0369 | -35.5849 | -37.7364
28.043 | -35.5852 | -37.7419
28.0491 | -35.5854 | -37.747
28.0552 | -35.5857 | -37.7519
28.0613 | -35.586 | -37.7569
28.0674 | -35.5863 | -37.7618
28.0735 | -35.5866 | -37.7665

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4575 | -42.1443
21.4908 | -35.4971 | -34.8913
16.4948 | -34.8841 | -27.5259
12.6602 | -34.5258 | -21.2705
9.7171 | -34.3152 | -16.3839
7.4582 | -34.1627 | -12.428
5.7244 | -34.0617 | -9.2171
4.3936 | -33.9906 | -6.6296
3.3722 | -33.9729 | -4.5911
2.5883 | -33.9634 | -3.0457
1.9866 | -33.9601 | -1.8106
1.5248 | -33.9644 | -0.82087
1.1703 | -33.9731 | -0.17359
1 | -33.9756 | 0.14596
0.89822 | -33.9639 | 0.41819
0.68941 | -33.9558 | 0.99818
0.52919 | -33.9628 | 1.4484
0.40612 | -33.9576 | 2.0537
0.31171 | -33.9571 | 2.5099
0.23928 | -33.975 | 3.2318
0.18358 | -33.9782 | 3.6529
0.14091 | -33.9733 | 4.5393
0.10819 | -34.0152 | 5.6884
0.083017 | -34.0187 | 7.1715
0.063705 | -34.077 | 9.0355
0.048923 | -34.1478 | 11.525
0.037479 | -34.2508 | 14.7024
0.028801 | -34.4309 | 18.707
0.022078 | -34.7303 | 24.0625
0.017023 | -35.2011 | 30.4998
0.013018 | -36.0001 | 38.2023
0.010014 | -37.1175 | 46.1097

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003751 | -8.3593
0.007501 | -26.5713
0.011252 | -33.3538
0.015002 | -36.8905
0.018753 | -40.0723
0.022504 | -42.8669
0.026254 | -44.1446
0.030005 | -46.4649
0.033755 | -49.0722
0.037506 | -50.1024
0.041256 | -51.7936
0.045007 | -53.1524
0.048758 | -53.8803
0.052508 | -53.5143
0.056259 | -54.6817
0.060009 | -56.7083
0.06376 | -57.993
0.067511 | -58.2714
0.071261 | -59.5497
0.075012 | -60.5196
0.078762 | -60.8681
0.082513 | -61.5938
0.086263 | -62.6033
0.090014 | -61.9963
0.093765 | -62.4928
0.097515 | -63.7987
0.10127 | -62.5305
0.10502 | -62.9571
0.10877 | -64.0232
0.11252 | -63.781
0.11627 | -64.0857
0.12002 | -65.1336
0.12377 | -65.6364
0.12752 | -66.2889
0.13127 | -65.9923
0.13502 | -66.2129
0.13877 | -66.6004
0.14252 | -67.5735
0.14627 | -68.1009
0.15002 | -69.2852
0.15377 | -69.3514
0.15752 | -68.8008
0.16128 | -69.1403
0.16503 | -69.3741
0.16878 | -70.227
0.17253 | -70.2699
0.17628 | -70.5281
0.18003 | -71.4687
0.18378 | -71.5468
0.18753 | -71.3181
0.19128 | -71.3048
0.19503 | -72.2983
0.19878 | -71.85
0.20253 | -72.1881
0.20628 | -72.1517
0.21003 | -72.4381
0.21378 | -72.4956
0.21753 | -72.5696
0.22129 | -73.4109
0.22504 | -72.948
0.22879 | -72.5317
0.23254 | -73.7026
0.23629 | -73.6358
0.24004 | -73.6709
0.24379 | -74.0063
0.24754 | -74.3634
0.25129 | -74.9694
0.25504 | -74.6157
0.25879 | -73.9808
0.26254 | -74.9172
0.26629 | -74.3563
0.27004 | -74.6178
0.27379 | -76.0361
0.27754 | -75.464
0.28129 | -75.0015
0.28504 | -75.9715
0.2888 | -75.6731
0.29255 | -75.168
0.2963 | -74.7496
0.30005 | -74.6698
0.3038 | -74.7739
0.30755 | -75.5132
0.3113 | -77.9363
0.31505 | -78.1084
0.3188 | -76.7618
0.32255 | -76.1005
0.3263 | -76.1963
0.33005 | -76.3802
0.3338 | -76.5471
0.33755 | -77.7862
0.3413 | -77.469
0.34505 | -77.0644
0.3488 | -76.8869
0.35256 | -76.6047
0.35631 | -76.7524
0.36006 | -77.7462
0.36381 | -78.1809
0.36756 | -78.3992
0.37131 | -78.3191
0.37506 | -78.1544
0.37881 | -77.923
0.38256 | -77.5807
0.38631 | -77.3843
0.39006 | -77.8794
0.39381 | -79.0073
0.39756 | -77.9582
0.40131 | -78.6117
0.40506 | -78.5456
0.40881 | -79.2491
0.41256 | -80.0205
0.41631 | -79.4887
0.42007 | -79.3524
0.42382 | -78.6806
0.42757 | -79.2506
0.43132 | -79.1864
0.43507 | -79.0865
0.43882 | -79.3427
0.44257 | -79.1673
0.44632 | -79.0097
0.45007 | -78.8543
0.45382 | -79.9677
0.45757 | -79.9295
0.46132 | -79.2644
0.46507 | -79.1145
0.46882 | -79.3646
0.47257 | -79.8531
0.47632 | -79.2616
0.48007 | -79.4322
0.48383 | -79.0017
0.48758 | -80.0965
0.49133 | -80.739
0.49508 | -80.3909
0.49883 | -80.7256
0.50258 | -80.3594
0.50633 | -81.1211
0.51008 | -81.0845
0.51383 | -80.9477
0.51758 | -80.8175
0.52133 | -81.5557
0.52508 | -82.045
0.52883 | -80.6402
0.53258 | -80.733
0.53633 | -80.9709
0.54008 | -81.2799
0.54383 | -82.027
0.54759 | -82.6121
0.55134 | -82.6587
0.55509 | -82.7088
0.55884 | -82.8378
0.56259 | -82.7248
0.56634 | -82.2401
0.57009 | -81.8624
0.57384 | -81.6646
0.57759 | -81.4781
0.58134 | -81.5998
0.58509 | -82.0666
0.58884 | -82.1922
0.59259 | -82.217
0.59634 | -82.2876
0.60009 | -82.1952
0.60384 | -82.2075
0.60759 | -82.4981
0.61135 | -82.7492
0.6151 | -82.9137
0.61885 | -82.6758
0.6226 | -82.2965
0.62635 | -82.3484
0.6301 | -82.5718
0.63385 | -82.758
0.6376 | -83.0743
0.64135 | -83.4383
0.6451 | -83.4545
0.64885 | -83.1702
0.6526 | -83.3349
0.65635 | -83.8977
0.6601 | -84.0707
0.66385 | -83.748
0.6676 | -83.2975
0.67136 | -83.0351
0.6751 | -82.9697
0.67886 | -82.9347
0.68261 | -82.8281
0.68636 | -83.1893
0.69011 | -83.7861
0.69386 | -84.0421
0.69761 | -84.0954
0.70136 | -83.7993
0.70511 | -83.7191
0.70886 | -83.6184
0.71261 | -83.5287
0.71636 | -83.7738
0.72011 | -83.9568
0.72386 | -84.1922
0.72761 | -84.5831
0.73136 | -84.5431
0.73511 | -84.43
0.73886 | -84.7689
0.74262 | -85.2501
0.74637 | -85.3686
0.75012 | -85.1653
0.75387 | -84.7052
0.75762 | -84.1632
0.76137 | -83.8903
0.76512 | -83.8291
0.76887 | -84.1025
0.77262 | -84.8151
0.77637 | -85.3722
0.78012 | -85.4738
0.78387 | -85.2222
0.78762 | -85.0979
0.79137 | -85.2075
0.79512 | -85.3628
0.79888 | -85.3614
0.80263 | -84.871
0.80638 | -84.7352
0.81013 | -85.1852
0.81388 | -85.445
0.81763 | -85.0347
0.82138 | -84.3325
0.82513 | -84.1497
0.82888 | -84.8402
0.83263 | -85.7382
0.83638 | -86.2183
0.84013 | -86.1151
0.84388 | -85.8943
0.84763 | -85.9933
0.85138 | -86.4709
0.85513 | -86.926
0.85888 | -86.8035
0.86264 | -86.7693
0.86638 | -86.3121
0.87014 | -85.6358
0.87389 | -85.316
0.87764 | -85.6925
0.88139 | -86.0204
0.88514 | -85.7345
0.88889 | -85.5817
0.89264 | -85.8546
0.89639 | -85.8212
0.90014 | -85.5501
0.90389 | -85.9084
0.90764 | -86.6905
0.91139 | -86.7461
0.91514 | -86.4486
0.91889 | -86.0045
0.92264 | -85.3416
0.92639 | -85.0261
0.93014 | -85.2628
0.9339 | -85.8931
0.93765 | -87.0686
0.9414 | -87.8167
0.94515 | -87.3024
0.9489 | -87.0649
0.95265 | -87.5796
0.9564 | -87.858
0.96015 | -87.4968
0.9639 | -87.1755
0.96765 | -87.1384
0.9714 | -86.8868
0.97515 | -86.8267
0.9789 | -87.3055
0.98265 | -87.8883
0.9864 | -87.8353
0.99016 | -87.5842
0.99391 | -86.8618
0.99766 | -85.9298
1.0014 | -85.9432
1.0052 | -86.7611
1.0089 | -87.2456
1.0127 | -86.946
1.0164 | -87.1123
1.0202 | -87.6154
1.0239 | -87.416
1.0277 | -87.2448
1.0314 | -87.6019
1.0352 | -88.035
1.0389 | -88.2127
1.0427 | -88.408
1.0464 | -88.0673
1.0502 | -87.6782
1.0539 | -87.4146
1.0577 | -87.5989
1.0614 | -88.2743
1.0652 | -88.9953
1.0689 | -89.5762
1.0727 | -89.5266
1.0764 | -88.7923
1.0802 | -87.8405
1.0839 | -87.4291
1.0877 | -87.563
1.0914 | -87.9813
1.0952 | -88.2659
1.0989 | -88.3363
1.1027 | -88.3048
1.1064 | -88.1453
1.1102 | -87.9753
1.1139 | -87.9966
1.1177 | -88.0923
1.1214 | -88.1328
1.1252 | -88.2682
1.1289 | -88.5351
1.1327 | -88.7837
1.1364 | -88.8708
1.1402 | -88.8477
1.1439 | -88.8972
1.1477 | -89.0188
1.1514 | -89.1555
1.1552 | -89.2749
1.1589 | -89.2282
1.1627 | -88.9939
1.1664 | -88.7501
1.1702 | -88.568
1.1739 | -88.4253
1.1777 | -88.4161
1.1814 | -88.635
1.1852 | -88.949
1.1889 | -89.0918
1.1927 | -88.9841
1.1964 | -88.7664
1.2002 | -88.575
1.2039 | -88.5057
1.2077 | -88.5465
1.2114 | -88.569
1.2152 | -88.5426
1.2189 | -88.5572
1.2227 | -88.7226
1.2264 | -89.0189
1.2302 | -89.2708
1.2339 | -89.3703
1.2377 | -89.3467
1.2414 | -89.2915
1.2452 | -89.3014
1.2489 | -89.3346
1.2527 | -89.2379
1.2564 | -89.0032
1.2602 | -88.7293
1.2639 | -88.5437
1.2677 | -88.58
1.2714 | -88.834
1.2752 | -89.2149
1.279 | -89.6601
1.2827 | -90.0252
1.2865 | -90.1337
1.2902 | -90.0452
1.294 | -89.9211
1.2977 | -89.7722
1.3015 | -89.6128
1.3052 | -89.5178
1.309 | -89.5462
1.3127 | -89.6457
1.3165 | -89.6071
1.3202 | -89.3638
1.324 | -89.1722
1.3277 | -89.3024
1.3315 | -89.7424
1.3352 | -90.217
1.339 | -90.472
1.3427 | -90.5221
1.3465 | -90.5762
1.3502 | -90.6018
1.354 | -90.3769
1.3577 | -90.0936
1.3615 | -90.0823
1.3652 | -90.3287
1.369 | -90.4863
1.3727 | -90.2219
1.3765 | -89.6803
1.3802 | -89.1411
1.384 | -88.8316
1.3877 | -88.9785
1.3915 | -89.4604
1.3952 | -89.9508
1.399 | -90.3642
1.4027 | -90.6058
1.4065 | -90.6047
1.4102 | -90.5854
1.414 | -90.6194
1.4177 | -90.5624
1.4215 | -90.4204
1.4252 | -90.3515
1.429 | -90.474
1.4327 | -90.6934
1.4365 | -90.8086
1.4402 | -90.8853
1.444 | -91.0503
1.4477 | -91.2126
1.4515 | -91.2441
1.4552 | -91.1332
1.459 | -90.9869
1.4627 | -90.9377
1.4665 | -91.0207
1.4702 | -91.0475
1.474 | -90.8674
1.4777 | -90.6499
1.4815 | -90.492
1.4852 | -90.3272
1.489 | -90.2436
1.4927 | -90.272
1.4965 | -90.1947
1.5002 | -90.0765
1.504 | -90.1991
1.5077 | -90.5095
1.5115 | -90.7865
1.5152 | -90.9901
1.519 | -91.0204
1.5227 | -90.7919
1.5265 | -90.5511
1.5302 | -90.6167
1.534 | -91.046
1.5377 | -91.5759
1.5415 | -91.7255
1.5452 | -91.3674
1.549 | -90.8859
1.5527 | -90.6166
1.5565 | -90.5758
1.5602 | -90.6556
1.564 | -90.8323
1.5677 | -91.1106
1.5715 | -91.4032
1.5752 | -91.5532
1.579 | -91.5233
1.5827 | -91.4258
1.5865 | -91.3493
1.5902 | -91.3607
1.594 | -91.3939
1.5978 | -91.2662
1.6015 | -91.107
1.6053 | -91.1647
1.609 | -91.4049
1.6128 | -91.6647
1.6165 | -91.7169
1.6203 | -91.6073
1.624 | -91.6512
1.6278 | -91.8071
1.6315 | -91.9364
1.6353 | -91.903
1.639 | -91.7294
1.6428 | -91.4414
1.6465 | -91.0985
1.6503 | -90.8463
1.654 | -90.8222
1.6578 | -91.0112
1.6615 | -91.2794
1.6653 | -91.5036
1.669 | -91.6572
1.6728 | -91.7603
1.6765 | -91.8344
1.6803 | -91.892
1.684 | -91.9522
1.6878 | -92.0276
1.6915 | -92.128
1.6953 | -92.2623
1.699 | -92.419
1.7028 | -92.5565
1.7065 | -92.6301
1.7103 | -92.6477
1.714 | -92.6747
1.7178 | -92.7954
1.7215 | -93.0247
1.7253 | -93.3004
1.729 | -93.4881
1.7328 | -93.5124
1.7365 | -93.398
1.7403 | -93.2078
1.744 | -92.9792
1.7478 | -92.7895
1.7515 | -92.7053
1.7553 | -92.758
1.759 | -92.8815
1.7628 | -92.9347
1.7665 | -92.8086
1.7703 | -92.5935
1.774 | -92.4344
1.7778 | -92.3933
1.7815 | -92.4316
1.7853 | -92.5037
1.789 | -92.5846
1.7928 | -92.6819
1.7965 | -92.789
1.8003 | -92.8556
1.804 | -92.8365
1.8078 | -92.7812
1.8115 | -92.7541
1.8153 | -92.7664
1.819 | -92.7889
1.8228 | -92.8115
1.8265 | -92.824
1.8303 | -92.8056
1.834 | -92.7381
1.8378 | -92.5982
1.8415 | -92.4461
1.8453 | -92.4116
1.849 | -92.5483
1.8528 | -92.741
1.8565 | -92.8146
1.8603 | -92.7174
1.864 | -92.538
1.8678 | -92.3533
1.8715 | -92.2327
1.8753 | -92.2286
1.879 | -92.3414
1.8828 | -92.4987
1.8865 | -92.6279
1.8903 | -92.7337
1.894 | -92.8579
1.8978 | -92.9614
1.9015 | -92.9749
1.9053 | -92.9399
1.909 | -93.0118
1.9128 | -93.2842
1.9165 | -93.6789
1.9203 | -93.9688
1.9241 | -93.9969
1.9278 | -93.8764
1.9316 | -93.8214
1.9353 | -93.8606
1.9391 | -93.8497
1.9428 | -93.6559
1.9466 | -93.2906
1.9503 | -92.8871
1.9541 | -92.5657
1.9578 | -92.3506
1.9616 | -92.2538
1.9653 | -92.3463
1.9691 | -92.6675
1.9728 | -93.1296
1.9766 | -93.5745
1.9803 | -93.8504
1.9841 | -93.9004
1.9878 | -93.8014
1.9916 | -93.6755
1.9953 | -93.5991
1.9991 | -93.5915
2.0028 | -93.6658
2.0066 | -93.8123
2.0103 | -93.9306
2.0141 | -93.9456
2.0178 | -93.8413
2.0216 | -93.6883
2.0253 | -93.6197
2.0291 | -93.6991
2.0328 | -93.8729
2.0366 | -93.9916
2.0403 | -93.931
2.0441 | -93.7428
2.0478 | -93.5403
2.0516 | -93.4501
2.0553 | -93.5072
2.0591 | -93.6005
2.0628 | -93.5906
2.0666 | -93.4423
2.0703 | -93.2573
2.0741 | -93.1643
2.0778 | -93.2669
2.0816 | -93.6056
2.0853 | -94.0551
2.0891 | -94.3739
2.0928 | -94.4311
2.0966 | -94.3008
2.1003 | -94.0924
2.1041 | -93.9253
2.1078 | -93.8876
2.1116 | -93.9833
2.1153 | -94.1398
2.1191 | -94.2926
2.1228 | -94.4091
2.1266 | -94.4606
2.1303 | -94.4006
2.1341 | -94.2646
2.1378 | -94.1252
2.1416 | -94.007
2.1453 | -93.8875
2.1491 | -93.788
2.1528 | -93.7568
2.1566 | -93.821
2.1603 | -93.9694
2.1641 | -94.1114
2.1678 | -94.1821
2.1716 | -94.1508
2.1753 | -94.0082
2.1791 | -93.8222
2.1828 | -93.6775
2.1866 | -93.6373
2.1903 | -93.7058
2.1941 | -93.8363
2.1978 | -93.9257
2.2016 | -93.8909
2.2053 | -93.7574
2.2091 | -93.6224
2.2128 | -93.5659
2.2166 | -93.6068
2.2203 | -93.6976
2.2241 | -93.7767
2.2278 | -93.8167
2.2316 | -93.8348
2.2353 | -93.8689
2.2391 | -93.9735
2.2428 | -94.1893
2.2466 | -94.509
2.2504 | -94.8682
2.2541 | -95.1734
2.2579 | -95.3073
2.2616 | -95.2514
2.2654 | -95.0948
2.2691 | -94.9403
2.2729 | -94.8647
2.2766 | -94.8582
2.2804 | -94.8355
2.2841 | -94.7178
2.2879 | -94.4836
2.2916 | -94.2139
2.2954 | -94.0283
2.2991 | -94.0172
2.3029 | -94.1705
2.3066 | -94.394
2.3104 | -94.5712
2.3141 | -94.6235
2.3179 | -94.574
2.3216 | -94.4944
2.3254 | -94.4457
2.3291 | -94.4843
2.3329 | -94.604
2.3366 | -94.7433
2.3404 | -94.8488
2.3441 | -94.8674
2.3479 | -94.7962
2.3516 | -94.6443
2.3554 | -94.4534
2.3591 | -94.2685
2.3629 | -94.1517
2.3666 | -94.1164
2.3704 | -94.1519
2.3741 | -94.2333
2.3779 | -94.3395
2.3816 | -94.4491
2.3854 | -94.5605
2.3891 | -94.6385
2.3929 | -94.6777
2.3966 | -94.6584
2.4004 | -94.5873
2.4041 | -94.4766
2.4079 | -94.3833
2.4116 | -94.3567
2.4154 | -94.4193
2.4191 | -94.5499
2.4229 | -94.6929
2.4266 | -94.7691
2.4304 | -94.7483
2.4341 | -94.6446
2.4379 | -94.4965
2.4416 | -94.3438
2.4454 | -94.2215
2.4491 | -94.1487
2.4529 | -94.1288
2.4566 | -94.1516
2.4604 | -94.2041
2.4641 | -94.2667
2.4679 | -94.3062
2.4716 | -94.2707
2.4754 | -94.1555
2.4791 | -94.0105
2.4829 | -93.8935
2.4866 | -93.8697
2.4904 | -93.9852
2.4941 | -94.216
2.4979 | -94.4956
2.5016 | -94.7269
2.5054 | -94.8725
2.5091 | -94.9394
2.5129 | -94.9832
2.5166 | -95.0268
2.5204 | -95.0869
2.5241 | -95.1567
2.5279 | -95.2107
2.5316 | -95.2342
2.5354 | -95.275
2.5391 | -95.3546
2.5429 | -95.4835
2.5466 | -95.6273
2.5504 | -95.7573
2.5541 | -95.8376
2.5579 | -95.8803
2.5617 | -95.8959
2.5654 | -95.8836
2.5692 | -95.8594
2.5729 | -95.8472
2.5767 | -95.8175
2.5804 | -95.7466
2.5842 | -95.6204
2.5879 | -95.4524
2.5917 | -95.2972
2.5954 | -95.2075
2.5992 | -95.2093
2.6029 | -95.2929
2.6067 | -95.4057
2.6104 | -95.4726
2.6142 | -95.4697
2.6179 | -95.4351
2.6217 | -95.3962
2.6254 | -95.3826
2.6292 | -95.4207
2.6329 | -95.5079
2.6367 | -95.6319
2.6404 | -95.7577
2.6442 | -95.865
2.6479 | -95.9673
2.6517 | -96.0883
2.6554 | -96.2211
2.6592 | -96.3388
2.6629 | -96.3837
2.6667 | -96.3198
2.6704 | -96.1982
2.6742 | -96.1125
2.6779 | -96.103
2.6817 | -96.133
2.6854 | -96.137
2.6892 | -96.0656
2.6929 | -95.9448
2.6967 | -95.8217
2.7004 | -95.7248
2.7042 | -95.747
2.7079 | -95.7639
2.7117 | -95.8312
2.7154 | -95.9365
2.7192 | -96.0618
2.7229 | -96.1831
2.7267 | -96.2771
2.7304 | -96.3239
2.7342 | -96.3298
2.7379 | -96.2937
2.7417 | -96.2327
2.7454 | -96.1758
2.7492 | -96.1452
2.7529 | -96.1414
2.7567 | -96.1519
2.7604 | -96.1573
2.7642 | -96.1443
2.7679 | -96.1102
2.7717 | -96.0587
2.7754 | -96.0072
2.7792 | -95.9896
2.7829 | -96.0388
2.7867 | -96.1579
2.7904 | -96.3123
2.7942 | -96.4654
2.7979 | -96.5545
2.8017 | -96.5698
2.8054 | -96.5127
2.8092 | -96.4144
2.8129 | -96.3152
2.8167 | -96.2662
2.8204 | -96.2948
2.8242 | -96.3825
2.8279 | -96.5006
2.8317 | -96.6016
2.8354 | -96.6675
2.8392 | -96.6997
2.8429 | -96.6941
2.8467 | -96.6893
2.8504 | -96.6693
2.8542 | -96.6383
2.8579 | -96.5899
2.8617 | -96.5158
2.8654 | -96.4101
2.8692 | -96.2819
2.8729 | -96.1577
2.8767 | -96.0667
2.8805 | -96.0611
2.8842 | -96.1543
2.888 | -96.331
2.8917 | -96.5458
2.8955 | -96.7442
2.8992 | -96.8941
2.903 | -96.9686
2.9067 | -96.981
2.9105 | -96.934
2.9142 | -96.8466
2.918 | -96.7521
2.9217 | -96.6639
2.9255 | -96.6077
2.9292 | -96.5835
2.933 | -96.585
2.9367 | -96.6098
2.9405 | -96.6498
2.9442 | -96.7048
2.948 | -96.7737
2.9517 | -96.8432
2.9555 | -96.8859
2.9592 | -96.8949
2.963 | -96.8418
2.9667 | -96.7414
2.9705 | -96.6271
2.9742 | -96.5168
2.978 | -96.4095
2.9817 | -96.3133
2.9855 | -96.22
2.9892 | -96.1358
2.993 | -96.0792
2.9967 | -96.0701
3.0005 | -96.1226
3.0042 | -96.2328
3.008 | -96.3741
3.0117 | -96.5217
3.0155 | -96.6295
3.0192 | -96.6786
3.023 | -96.6848
3.0267 | -96.7023
3.0305 | -96.7462
3.0342 | -96.8105
3.038 | -96.881
3.0417 | -96.9159
3.0455 | -96.8878
3.0492 | -96.7904
3.053 | -96.6688
3.0567 | -96.5764
3.0605 | -96.5665
3.0642 | -96.6666
3.068 | -96.8679
3.0717 | -97.1278
3.0755 | -97.3592
3.0792 | -97.4932
3.083 | -97.5102
3.0867 | -97.4118
3.0905 | -97.2493
3.0942 | -97.0721
3.098 | -96.9143
3.1017 | -96.79
3.1055 | -96.7006
3.1092 | -96.6347
3.113 | -96.5844
3.1167 | -96.5496
3.1205 | -96.548
3.1242 | -96.6038
3.128 | -96.7066
3.1317 | -96.8352
3.1355 | -96.9643
3.1392 | -97.0608
3.143 | -97.1243
3.1467 | -97.1201
3.1505 | -97.04
3.1542 | -96.9091
3.158 | -96.7673
3.1617 | -96.6772
3.1655 | -96.6857
3.1692 | -96.8022
3.173 | -97.0088
3.1767 | -97.232
3.1805 | -97.3799
3.1842 | -97.3983
3.188 | -97.2977
3.1917 | -97.1207
3.1955 | -96.9667
3.1993 | -96.8807
3.203 | -96.878
3.2068 | -96.9442
3.2105 | -97.0384
3.2143 | -97.1302
3.218 | -97.1883
3.2218 | -97.1979
3.2255 | -97.1795
3.2293 | -97.1443
3.233 | -97.1245
3.2368 | -97.1302
3.2405 | -97.1469
3.2443 | -97.0868
3.248 | -97.029
3.2518 | -96.9467
3.2555 | -96.861
3.2593 | -96.7955
3.263 | -96.7715
3.2668 | -96.7889
3.2705 | -96.8548
3.2743 | -96.9538
3.278 | -97.0803
3.2818 | -97.2216
3.2855 | -97.3448
3.2893 | -97.4512
3.293 | -97.5451
3.2968 | -97.6188
3.3005 | -97.6819
3.3043 | -97.7154
3.308 | -97.7233
3.3118 | -97.7001
3.3155 | -97.6479
3.3193 | -97.5859
3.323 | -97.5519
3.3268 | -97.5571
3.3305 | -97.6081
3.3343 | -97.6673
3.338 | -97.7211
3.3418 | -97.745
3.3455 | -97.711
3.3493 | -97.6513
3.353 | -97.5934
3.3568 | -97.5439
3.3605 | -97.5141
3.3643 | -97.4969
3.368 | -97.4998
3.3718 | -97.521
3.3755 | -97.5476
3.3793 | -97.5562
3.383 | -97.5348
3.3868 | -97.4825
3.3905 | -97.4253
3.3943 | -97.3739
3.398 | -97.3609
3.4018 | -97.3985
3.4055 | -97.4646
3.4093 | -97.5561
3.413 | -97.6567
3.4168 | -97.7547
3.4205 | -97.843
3.4243 | -97.9038
3.428 | -97.9309
3.4318 | -97.926
3.4355 | -97.8922
3.4393 | -97.8516
3.443 | -97.8175
3.4468 | -97.8029
3.4505 | -97.8232
3.4543 | -97.8775
3.458 | -97.9597
3.4618 | -98.0634
3.4655 | -98.1579
3.4693 | -98.2144
3.473 | -98.2294
3.4768 | -98.2028
3.4805 | -98.135
3.4843 | -98.0491
3.488 | -97.9537
3.4918 | -97.8638
3.4955 | -97.7968
3.4993 | -97.7591
3.503 | -97.7501
3.5068 | -97.7398
3.5105 | -97.7148
3.5143 | -97.6691
3.5181 | -97.6056
3.5218 | -97.546
3.5256 | -97.5035
3.5293 | -97.465
3.5331 | -97.4442
3.5368 | -97.46
3.5406 | -97.5103
3.5443 | -97.5994
3.5481 | -97.7202
3.5518 | -97.8468
3.5556 | -97.9758
3.5593 | -98.1041
3.5631 | -98.2354
3.5668 | -98.3501
3.5706 | -98.4402
3.5743 | -98.4867
3.5781 | -98.4869
3.5818 | -98.4419
3.5856 | -98.3617
3.5893 | -98.2619
3.5931 | -98.1673
3.5968 | -98.0931
3.6006 | -98.0586
3.6043 | -98.0753
3.6081 | -98.1375
3.6118 | -98.2191
3.6156 | -98.2983
3.6193 | -98.3641
3.6231 | -98.3899
3.6268 | -98.3681
3.6306 | -98.3245
3.6343 | -98.268
3.6381 | -98.2307
3.6418 | -98.2247
3.6456 | -98.2371
3.6493 | -98.2579
3.6531 | -98.2844
3.6568 | -98.3086
3.6606 | -98.3406
3.6643 | -98.3878
3.6681 | -98.4424
3.6718 | -98.4891
3.6756 | -98.5103
3.6793 | -98.5075
3.6831 | -98.4735
3.6868 | -98.4012
3.6906 | -98.3215
3.6943 | -98.2487
3.6981 | -98.2086
3.7018 | -98.2074
3.7056 | -98.2579
3.7093 | -98.3549
3.7131 | -98.4658
3.7168 | -98.5573
3.7206 | -98.607
3.7243 | -98.5995
3.7281 | -98.5322
3.7318 | -98.4166
3.7356 | -98.2823
3.7393 | -98.1683
3.7431 | -98.0897
3.7468 | -98.0405
3.7506 | -98.0159
3.7543 | -98.0077
3.7581 | -98.0221
3.7618 | -98.0586
3.7656 | -98.1203
3.7693 | -98.225
3.7731 | -98.3636
3.7768 | -98.5112
3.7806 | -98.6423
3.7843 | -98.6965
3.7881 | -98.7617
3.7918 | -98.8059
3.7956 | -98.835
3.7993 | -98.8459
3.8031 | -98.8573
3.8068 | -98.873
3.8106 | -98.8815
3.8143 | -98.8973
3.8181 | -98.9217
3.8218 | -98.9589
3.8256 | -99.0221
3.8293 | -99.0922
3.8331 | -99.1545
3.8369 | -99.1993
3.8406 | -99.2116
3.8444 | -99.1789
3.8481 | -99.1053
3.8519 | -99.0107
3.8556 | -98.9142
3.8594 | -98.8302
3.8631 | -98.7638
3.8669 | -98.7198
3.8706 | -98.6913
3.8744 | -98.6837
3.8781 | -98.6739
3.8819 | -98.6522
3.8856 | -98.6232
3.8894 | -98.5949
3.8931 | -98.5748
3.8969 | -98.5675
3.9006 | -98.5856
3.9044 | -98.6232
3.9081 | -98.6751
3.9119 | -98.7305
3.9156 | -98.7794
3.9194 | -98.8138
3.9231 | -98.8284
3.9269 | -98.8212
3.9306 | -98.7835
3.9344 | -98.7389
3.9381 | -98.6904
3.9419 | -98.6408
3.9456 | -98.5999
3.9494 | -98.5628
3.9531 | -98.5301
3.9569 | -98.5092
3.9606 | -98.5113
3.9644 | -98.5406
3.9681 | -98.6044
3.9719 | -98.7014
3.9756 | -98.8213
3.9794 | -98.9389
3.9831 | -99.0519
3.9869 | -99.1513
3.9906 | -99.2281
3.9944 | -99.2813
3.9981 | -99.2945
4.0019 | -99.2805
4.0056 | -99.2511
4.0094 | -99.2044
4.0131 | -99.1553
4.0169 | -99.1073
4.0206 | -99.0758
4.0244 | -99.048
4.0281 | -99.0174
4.0319 | -98.9859
4.0356 | -98.9569
4.0394 | -98.9431
4.0431 | -98.9427
4.0469 | -98.9379
4.0506 | -98.931
4.0544 | -98.9237
4.0581 | -98.9103
4.0619 | -98.903
4.0656 | -98.8996
4.0694 | -98.9164
4.0731 | -98.9495
4.0769 | -98.9772
4.0806 | -98.9886
4.0844 | -98.9764
4.0881 | -98.9482
4.0919 | -98.9176
4.0956 | -98.8943
4.0994 | -98.8791
4.1031 | -98.8794
4.1069 | -98.9067
4.1106 | -98.9554
4.1144 | -99.0055
4.1181 | -99.0585
4.1219 | -99.107
4.1256 | -99.1413
4.1294 | -99.1661
4.1331 | -99.1671
4.1369 | -99.1491
4.1406 | -99.1327
4.1444 | -99.1225
4.1481 | -99.1224
4.1519 | -99.1404
4.1556 | -99.1728
4.1594 | -99.2182
4.1631 | -99.2586
4.1669 | -99.2888
4.1707 | -99.3007
4.1744 | -99.2962
4.1782 | -99.2932
4.1819 | -99.2897
4.1857 | -99.2892
4.1894 | -99.2958
4.1932 | -99.3052
4.1969 | -99.326
4.2007 | -99.3624
4.2044 | -99.3992
4.2082 | -99.4314
4.2119 | -99.4629
4.2157 | -99.4911
4.2194 | -99.5181
4.2232 | -99.5447
4.2269 | -99.5741
4.2307 | -99.5961
4.2344 | -99.6046
4.2382 | -99.5825
4.2419 | -99.5259
4.2457 | -99.445
4.2494 | -99.3511
4.2532 | -99.2533
4.2569 | -99.1609
4.2607 | -99.0811
4.2644 | -99.0224
4.2682 | -98.9971
4.2719 | -99.0062
4.2757 | -99.0363
4.2794 | -99.0886
4.2832 | -99.1597
4.2869 | -99.2403
4.2907 | -99.3289
4.2944 | -99.4051
4.2982 | -99.4717
4.3019 | -99.5214
4.3057 | -99.5546
4.3094 | -99.576
4.3132 | -99.5961
4.3169 | -99.6155
4.3207 | -99.6484
4.3244 | -99.6912
4.3282 | -99.7795
4.3319 | -99.8477
4.3357 | -99.9216
4.3394 | -99.9877
4.3432 | -100.0311
4.3469 | -100.0409
4.3507 | -100.0102
4.3544 | -99.9417
4.3582 | -99.8406
4.3619 | -99.7339
4.3657 | -99.6334
4.3694 | -99.5511
4.3732 | -99.47
4.3769 | -99.4072
4.3807 | -99.3629
4.3844 | -99.3308
4.3882 | -99.3171
4.3919 | -99.3271
4.3957 | -99.3486
4.3994 | -99.3802
4.4032 | -99.4262
4.4069 | -99.4745
4.4107 | -99.5282
4.4144 | -99.5732
4.4182 | -99.6129
4.4219 | -99.6472
4.4257 | -99.6698
4.4294 | -99.672
4.4332 | -99.6444
4.4369 | -99.592
4.4407 | -99.5327
4.4444 | -99.4799
4.4482 | -99.4393
4.4519 | -99.4174
4.4557 | -99.4234
4.4594 | -99.4461
4.4632 | -99.4805
4.4669 | -99.5203
4.4707 | -99.5507
4.4744 | -99.5647
4.4782 | -99.5722
4.482 | -99.5746
4.4857 | -99.5823
4.4895 | -99.5928
4.4932 | -99.6097
4.497 | -99.6442
4.5007 | -99.6893
4.5045 | -99.7455
4.5082 | -99.8004
4.512 | -99.8497
4.5157 | -99.8973
4.5195 | -99.9341
4.5232 | -99.9654
4.527 | -99.9914
4.5307 | -100.0099
4.5345 | -100.016
4.5382 | -100.0091
4.542 | -99.9888
4.5457 | -99.9635
4.5495 | -99.9275
4.5532 | -99.8886
4.557 | -99.8592
4.5607 | -99.8363
4.5645 | -99.8269
4.5682 | -99.8262
4.572 | -99.8428
4.5757 | -99.8649
4.5795 | -99.8939
4.5832 | -99.921
4.587 | -99.9471
4.5907 | -99.9667
4.5945 | -99.97
4.5982 | -99.9607
4.602 | -99.9461
4.6057 | -99.9328
4.6095 | -99.9192
4.6132 | -99.9057
4.617 | -99.893
4.6207 | -99.8781
4.6245 | -99.8648
4.6282 | -99.8566
4.632 | -99.8591
4.6357 | -99.884
4.6395 | -99.9346
4.6432 | -100.0065
4.647 | -100.0971
4.6507 | -100.1968
4.6545 | -100.294
4.6582 | -100.3743
4.662 | -100.4322
4.6657 | -100.4623
4.6695 | -100.4595
4.6732 | -100.4206
4.677 | -100.3654
4.6807 | -100.3061
4.6845 | -100.2444
4.6882 | -100.1956
4.692 | -100.176
4.6957 | -100.1707
4.6995 | -100.1806
4.7032 | -100.2021
4.707 | -100.2203
4.7107 | -100.2327
4.7145 | -100.2325
4.7182 | -100.209
4.722 | -100.1647
4.7257 | -100.1072
4.7295 | -100.0438
4.7332 | -99.9923
4.737 | -99.9582
4.7407 | -99.9427
4.7445 | -99.9594
4.7482 | -99.9937
4.752 | -100.0442
4.7557 | -100.1155
4.7595 | -100.1896
4.7632 | -100.2592
4.767 | -100.3199
4.7707 | -100.3634
4.7745 | -100.3887
4.7782 | -100.3884
4.782 | -100.3667
4.7857 | -100.3386
4.7895 | -100.3128
4.7932 | -100.281
4.797 | -100.2471
4.8007 | -100.2136
4.8045 | -100.1861
4.8083 | -100.1602
4.812 | -100.1389
4.8158 | -100.1231
4.8195 | -100.1127
4.8233 | -100.1084
4.827 | -100.1096
4.8308 | -100.1196
4.8345 | -100.1385
4.8383 | -100.1748
4.842 | -100.2191
4.8458 | -100.2675
4.8495 | -100.3214
4.8533 | -100.3732
4.857 | -100.4148
4.8608 | -100.4439
4.8645 | -100.4767
4.8683 | -100.493
4.872 | -100.522
4.8758 | -100.5436
4.8795 | -100.5581
4.8833 | -100.5812
4.887 | -100.5989
4.8908 | -100.6139
4.8945 | -100.6363
4.8983 | -100.6602
4.902 | -100.6859
4.9058 | -100.717
4.9095 | -100.7508
4.9133 | -100.7711
4.917 | -100.7779
4.9208 | -100.7734
4.9245 | -100.7555
4.9283 | -100.7271
4.932 | -100.6998
4.9358 | -100.6791
4.9395 | -100.6617
4.9433 | -100.6641
4.947 | -100.6733
4.9508 | -100.6874
4.9545 | -100.7088
4.9583 | -100.7217
4.962 | -100.7165
4.9658 | -100.6992
4.9695 | -100.6658
4.9733 | -100.6266
4.977 | -100.5813
4.9808 | -100.5296
4.9845 | -100.4734
4.9883 | -100.4225
4.992 | -100.3853
4.9958 | -100.3608
4.9995 | -100.3555
5.0033 | -100.3694
5.007 | -100.406
5.0108 | -100.4534
5.0145 | -100.5161
5.0183 | -100.5861
5.022 | -100.6549
5.0258 | -100.7245
5.0295 | -100.7895
5.0333 | -100.8453
5.037 | -100.8881
5.0408 | -100.921
5.0445 | -100.9364
5.0483 | -100.9407
5.052 | -100.9274
5.0558 | -100.9068
5.0595 | -100.8806
5.0633 | -100.8478
5.067 | -100.8007
5.0708 | -100.7516
5.0745 | -100.7034
5.0783 | -100.6494
5.082 | -100.5952
5.0858 | -100.5428
5.0895 | -100.5127
5.0933 | -100.499
5.097 | -100.4994
5.1008 | -100.5271
5.1045 | -100.5774
5.1083 | -100.6406
5.112 | -100.7137
5.1158 | -100.7872
5.1196 | -100.8392
5.1233 | -100.8711
5.1271 | -100.8753
5.1308 | -100.8627
5.1346 | -100.8322
5.1383 | -100.7944
5.1421 | -100.7559
5.1458 | -100.7199
5.1496 | -100.6856
5.1533 | -100.6589
5.1571 | -100.6424
5.1608 | -100.6383
5.1646 | -100.6406
5.1683 | -100.6417
5.1721 | -100.652
5.1758 | -100.6712
5.1796 | -100.6898
5.1833 | -100.7029
5.1871 | -100.7146
5.1908 | -100.7245
5.1946 | -100.7307
5.1983 | -100.732
5.2021 | -100.7224
5.2058 | -100.7069
5.2096 | -100.6836
5.2133 | -100.658
5.2171 | -100.6322
5.2208 | -100.6035
5.2246 | -100.5825
5.2283 | -100.5752
5.2321 | -100.5711
5.2358 | -100.5816
5.2396 | -100.6084
5.2433 | -100.6566
5.2471 | -100.728
5.2508 | -100.8169
5.2546 | -100.9235
5.2583 | -101.0328
5.2621 | -101.1499
5.2658 | -101.2546
5.2696 | -101.3424
5.2733 | -101.4085
5.2771 | -101.4496
5.2808 | -101.4642
5.2846 | -101.4507
5.2883 | -101.42
5.2921 | -101.3791
5.2958 | -101.3341
5.2996 | -101.2914
5.3033 | -101.2454
5.3071 | -101.1995
5.3108 | -101.1686
5.3146 | -101.1351
5.3183 | -101.0993
5.3221 | -101.0601
5.3258 | -101.0152
5.3296 | -100.971
5.3333 | -100.9313
5.3371 | -100.8994
5.3408 | -100.8659
5.3446 | -100.8404
5.3483 | -100.818
5.3521 | -100.8016
5.3558 | -100.7908
5.3596 | -100.7796
5.3633 | -100.7702
5.3671 | -100.7686
5.3708 | -100.7729
5.3746 | -100.787
5.3783 | -100.8207
5.3821 | -100.8775
5.3858 | -100.9507
5.3896 | -101.0321
5.3933 | -101.1158
5.3971 | -101.1956
5.4008 | -101.2732
5.4046 | -101.3337
5.4083 | -101.3354
5.4121 | -101.353
5.4158 | -101.3566
5.4196 | -101.3528
5.4233 | -101.347
5.4271 | -101.337
5.4308 | -101.3299
5.4346 | -101.3387
5.4383 | -101.3656
5.4421 | -101.4167
5.4459 | -101.4805
5.4496 | -101.5446
5.4534 | -101.6122
5.4571 | -101.6776
5.4609 | -101.735
5.4646 | -101.7749
5.4684 | -101.797
5.4721 | -101.8135
5.4759 | -101.8289
5.4796 | -101.8294
5.4834 | -101.8094
5.4871 | -101.7841
5.4909 | -101.7514
5.4946 | -101.7169
5.4984 | -101.6862
5.5021 | -101.6465
5.5059 | -101.6029
5.5096 | -101.5559
5.5134 | -101.5081
5.5171 | -101.4652
5.5209 | -101.4244
5.5246 | -101.3809
5.5284 | -101.3356
5.5321 | -101.3004
5.5359 | -101.2725
5.5396 | -101.2427
5.5434 | -101.2095
5.5471 | -101.1806
5.5509 | -101.1582
5.5546 | -101.14
5.5584 | -101.1286
5.5621 | -101.1267
5.5659 | -101.1358
5.5696 | -101.1606
5.5734 | -101.1881
5.5771 | -101.2204
5.5809 | -101.2545
5.5846 | -101.2845
5.5884 | -101.3104
5.5921 | -101.3354
5.5959 | -101.3642
5.5996 | -101.3882
5.6034 | -101.4029
5.6071 | -101.4121
5.6109 | -101.4144
5.6146 | -101.4117
5.6184 | -101.4003
5.6221 | -101.3862
5.6259 | -101.3763
5.6296 | -101.3777
5.6334 | -101.3868
5.6371 | -101.4033
5.6409 | -101.4297
5.6446 | -101.4547
5.6484 | -101.4788
5.6521 | -101.5055
5.6559 | -101.5217
5.6596 | -101.5415
5.6634 | -101.5494
5.6671 | -101.5577
5.6709 | -101.5666
5.6746 | -101.5768
5.6784 | -101.5836
5.6821 | -101.5807
5.6859 | -101.5754
5.6896 | -101.5708
5.6934 | -101.5563
5.6971 | -101.5403
5.7009 | -101.5239
5.7046 | -101.5058
5.7084 | -101.4875
5.7121 | -101.4614
5.7159 | -101.4317
5.7196 | -101.409
5.7234 | -101.3995
5.7271 | -101.3991
5.7309 | -101.4073
5.7346 | -101.4277
5.7384 | -101.4695
5.7421 | -101.5357
5.7459 | -101.6068
5.7496 | -101.6935
5.7534 | -101.7828
5.7572 | -101.8672
5.7609 | -101.9461
5.7647 | -102.0166
5.7684 | -102.0835
5.7722 | -102.1379
5.7759 | -102.1895
5.7797 | -102.2407
5.7834 | -102.2811
5.7872 | -102.293
5.7909 | -102.2915
5.7947 | -102.27
5.7984 | -102.2376
5.8022 | -102.1878
5.8059 | -102.121
5.8097 | -102.0521
5.8134 | -101.9915
5.8172 | -101.9345
5.8209 | -101.8832
5.8247 | -101.8343
5.8284 | -101.79
5.8322 | -101.7628
5.8359 | -101.7463
5.8397 | -101.7439
5.8434 | -101.7589
5.8472 | -101.7783
5.8509 | -101.8009
5.8547 | -101.8209
5.8584 | -101.8343
5.8622 | -101.8313
5.8659 | -101.8222
5.8697 | -101.8
5.8734 | -101.7816
5.8772 | -101.7561
5.8809 | -101.7423
5.8847 | -101.7358
5.8884 | -101.7436
5.8922 | -101.7672
5.8959 | -101.8009
5.8997 | -101.8469
5.9034 | -101.9053
5.9072 | -101.9597
5.9109 | -102.0122
5.9147 | -102.0527
5.9184 | -102.0712
5.9222 | -102.0733
5.9259 | -102.046
5.9297 | -101.9962
5.9334 | -101.9402
5.9372 | -101.8713
5.9409 | -101.8018
5.9447 | -101.7361
5.9484 | -101.699
5.9522 | -101.6524
5.9559 | -101.6131
5.9597 | -101.5943
5.9634 | -101.5762
5.9672 | -101.5676
5.9709 | -101.5673
5.9747 | -101.5736
5.9784 | -101.5746
5.9822 | -101.5737
5.9859 | -101.5664
5.9897 | -101.554
5.9934 | -101.5369
5.9972 | -101.5183
6.0009 | -101.4964
6.0047 | -101.478
6.0084 | -101.4618
6.0122 | -101.4492
6.0159 | -101.4519
6.0197 | -101.4637
6.0234 | -101.4814
6.0272 | -101.5169
6.0309 | -101.566
6.0347 | -101.6248
6.0384 | -101.6886
6.0422 | -101.7631
6.0459 | -101.8442
6.0497 | -101.9281
6.0534 | -102.0078
6.0572 | -102.0912
6.0609 | -102.1698
6.0647 | -102.241
6.0684 | -102.3089
6.0722 | -102.3718
6.0759 | -102.4246
6.0797 | -102.473
6.0835 | -102.5079
6.0872 | -102.5256
6.091 | -102.531
6.0947 | -102.5234
6.0985 | -102.509
6.1022 | -102.4914
6.106 | -102.4647
6.1097 | -102.4247
6.1135 | -102.3775
6.1172 | -102.327
6.121 | -102.2689
6.1247 | -102.2069
6.1285 | -102.1546
6.1322 | -102.1134
6.136 | -102.0837
6.1397 | -102.0688
6.1435 | -102.0656
6.1472 | -102.0802
6.151 | -102.1048
6.1547 | -102.1365
6.1585 | -102.1707
6.1622 | -102.2087
6.166 | -102.2543
6.1697 | -102.2995
6.1735 | -102.3427
6.1772 | -102.3881
6.181 | -102.4359
6.1847 | -102.4743
6.1885 | -102.5087
6.1922 | -102.5359
6.196 | -102.5477
6.1997 | -102.5535
6.2035 | -102.5567
6.2072 | -102.5515
6.211 | -102.5341
6.2147 | -102.5091
6.2185 | -102.4777
6.2222 | -102.4323
6.226 | -102.3785
6.2297 | -102.3176
6.2335 | -102.2575
6.2372 | -102.2046
6.241 | -102.1588
6.2447 | -102.1275
6.2485 | -102.1094
6.2522 | -102.0981
6.256 | -102.0972
6.2597 | -102.0974
6.2635 | -102.1055
6.2672 | -102.1161
6.271 | -102.1193
6.2747 | -102.1208
6.2785 | -102.1224
6.2822 | -102.1145
6.286 | -102.1027
6.2897 | -102.096
6.2935 | -102.0877
6.2972 | -102.0818
6.301 | -102.0813
6.3047 | -102.0876
6.3085 | -102.101
6.3122 | -102.1219
6.316 | -102.1542
6.3197 | -102.1928
6.3235 | -102.2373
6.3272 | -102.2946
6.331 | -102.3577
6.3347 | -102.4258
6.3385 | -102.4889
6.3422 | -102.5455
6.346 | -102.6068
6.3497 | -102.6525
6.3535 | -102.6922
6.3572 | -102.7043
6.361 | -102.702
6.3647 | -102.6962
6.3685 | -102.6744
6.3722 | -102.6415
6.376 | -102.6041
6.3797 | -102.5664
6.3835 | -102.5335
6.3872 | -102.5028
6.391 | -102.4769
6.3947 | -102.4566
6.3985 | -102.4452
6.4023 | -102.4421
6.406 | -102.4422
6.4098 | -102.4428
6.4135 | -102.448
6.4173 | -102.4498
6.421 | -102.4436
6.4248 | -102.4374
6.4285 | -102.4299
6.4323 | -102.4306
6.436 | -102.4398
6.4398 | -102.4489
6.4435 | -102.4687
6.4473 | -102.4958
6.451 | -102.5387
6.4548 | -102.5821
6.4585 | -102.6209
6.4623 | -102.666
6.466 | -102.7177
6.4698 | -102.7722
6.4735 | -102.8258
6.4773 | -102.8745
6.481 | -102.9219
6.4848 | -102.9611
6.4885 | -102.9851
6.4923 | -102.9581
6.496 | -102.9453
6.4998 | -102.9202
6.5035 | -102.8891
6.5073 | -102.8499
6.511 | -102.8016
6.5148 | -102.7596
6.5185 | -102.7188
6.5223 | -102.677
6.526 | -102.6359
6.5298 | -102.6013
6.5335 | -102.5685
6.5373 | -102.5454
6.541 | -102.5284
6.5448 | -102.5144
6.5485 | -102.5056
6.5523 | -102.5031
6.556 | -102.4977
6.5598 | -102.4847
6.5635 | -102.4695
6.5673 | -102.4589
6.571 | -102.4539
6.5748 | -102.4517
6.5785 | -102.4533
6.5823 | -102.4669
6.586 | -102.4925
6.5898 | -102.5222
6.5935 | -102.5608
6.5973 | -102.6072
6.601 | -102.656
6.6048 | -102.7063
6.6085 | -102.7608
6.6123 | -102.8147
6.616 | -102.8589
6.6198 | -102.8967
6.6235 | -102.9382
6.6273 | -102.9737
6.631 | -102.9979
6.6348 | -103.0212
6.6385 | -103.0345
6.6423 | -103.0474
6.646 | -103.0631
6.6498 | -103.0783
6.6535 | -103.0931
6.6573 | -103.1179
6.661 | -103.14
6.6648 | -103.1546
6.6685 | -103.1679
6.6723 | -103.1758
6.676 | -103.1802
6.6798 | -103.1844
6.6835 | -103.1842
6.6873 | -103.1753
6.691 | -103.1691
6.6948 | -103.1502
6.6985 | -103.1293
6.7023 | -103.0953
6.706 | -103.061
6.7098 | -103.0235
6.7135 | -102.981
6.7173 | -102.9392
6.721 | -102.897
6.7248 | -102.8414
6.7286 | -102.7869
6.7323 | -102.7303
6.7361 | -102.6734
6.7398 | -102.6274
6.7436 | -102.5858
6.7473 | -102.5454
6.7511 | -102.5148
6.7548 | -102.4901
6.7586 | -102.4694
6.7623 | -102.4593
6.7661 | -102.4567
6.7698 | -102.4568
6.7736 | -102.4619
6.7773 | -102.4762
6.7811 | -102.4893
6.7848 | -102.4948
6.7886 | -102.502
6.7923 | -102.5081
6.7961 | -102.5199
6.7998 | -102.529
6.8036 | -102.5426
6.8073 | -102.5567
6.8111 | -102.5745
6.8148 | -102.5995
6.8186 | -102.6301
6.8223 | -102.6697
6.8261 | -102.7202
6.8298 | -102.7666
6.8336 | -102.8083
6.8373 | -102.8487
6.8411 | -102.8859
6.8448 | -102.9224
6.8486 | -102.9564
6.8523 | -102.9814
6.8561 | -103.0014
6.8598 | -103.0155
6.8636 | -103.0175
6.8673 | -103.005
6.8711 | -102.9821
6.8748 | -102.9595
6.8786 | -102.9317
6.8823 | -102.9004
6.8861 | -102.8689
6.8898 | -102.8417
6.8936 | -102.8156
6.8973 | -102.7909
6.9011 | -102.7707
6.9048 | -102.7558
6.9086 | -102.7412
6.9123 | -102.7302
6.9161 | -102.7268
6.9198 | -102.7259
6.9236 | -102.726
6.9273 | -102.7311
6.9311 | -102.7451
6.9348 | -102.7613
6.9386 | -102.7768
6.9423 | -102.7912
6.9461 | -102.8103
6.9498 | -102.8353
6.9536 | -102.8652
6.9573 | -102.8926
6.9611 | -102.9199
6.9648 | -102.9541
6.9686 | -102.9933
6.9723 | -103.0243
6.9761 | -103.046
6.9798 | -103.0629
6.9836 | -103.077
6.9873 | -103.085
6.9911 | -103.0809
6.9948 | -103.0762
6.9986 | -103.0703
7.0023 | -103.0602
7.0061 | -103.0539
7.0098 | -103.0476
7.0136 | -103.0396
7.0173 | -103.0308
7.0211 | -103.0239
7.0248 | -103.0229
7.0286 | -103.0249
7.0323 | -103.041
7.0361 | -103.051
7.0399 | -103.0659
7.0436 | -103.0781
7.0474 | -103.0958
7.0511 | -103.1183
7.0549 | -103.1457
7.0586 | -103.1779
7.0624 | -103.2111
7.0661 | -103.2402
7.0699 | -103.2628
7.0736 | -103.2771
7.0774 | -103.2951
7.0811 | -103.3157
7.0849 | -103.3352
7.0886 | -103.3526
7.0924 | -103.3749
7.0961 | -103.3959
7.0999 | -103.4127
7.1036 | -103.4294
7.1074 | -103.4435
7.1111 | -103.4462
7.1149 | -103.4414
7.1186 | -103.4279
7.1224 | -103.4041
7.1261 | -103.3709
7.1299 | -103.3255
7.1336 | -103.2773
7.1374 | -103.2295
7.1411 | -103.178
7.1449 | -103.1294
7.1486 | -103.0853
7.1524 | -103.0505
7.1561 | -103.0214
7.1599 | -102.9985
7.1636 | -102.99
7.1674 | -102.9942
7.1711 | -103.0105
7.1749 | -103.0363
7.1786 | -103.0707
7.1824 | -103.1114
7.1861 | -103.1538
7.1899 | -103.1991
7.1936 | -103.2504
7.1974 | -103.3057
7.2011 | -103.3539
7.2049 | -103.3958
7.2086 | -103.4315
7.2124 | -103.4657
7.2161 | -103.497
7.2199 | -103.5233
7.2236 | -103.5468
7.2274 | -103.566
7.2311 | -103.5791
7.2349 | -103.592
7.2386 | -103.6
7.2424 | -103.6037
7.2461 | -103.6036
7.2499 | -103.6068
7.2536 | -103.6064
7.2574 | -103.6058
7.2611 | -103.6033
7.2649 | -103.6013
7.2686 | -103.5964
7.2724 | -103.5866
7.2761 | -103.5864
7.2799 | -103.59
7.2836 | -103.5974
7.2874 | -103.6102
7.2911 | -103.627
7.2949 | -103.6409
7.2986 | -103.6471
7.3024 | -103.6481
7.3061 | -103.6424
7.3099 | -103.6347
7.3136 | -103.6209
7.3174 | -103.5986
7.3211 | -103.5661
7.3249 | -103.5277
7.3286 | -103.4863
7.3324 | -103.4434
7.3361 | -103.3989
7.3399 | -103.359
7.3436 | -103.3232
7.3474 | -103.2997
7.3511 | -103.2799
7.3549 | -103.266
7.3586 | -103.2605
7.3624 | -103.2658
7.3662 | -103.2765
7.3699 | -103.2931
7.3737 | -103.3121
7.3774 | -103.3401
7.3812 | -103.3723
7.3849 | -103.3968
7.3887 | -103.4191
7.3924 | -103.4302
7.3962 | -103.4345
7.3999 | -103.4348
7.4037 | -103.4319
7.4074 | -103.42
7.4112 | -103.3999
7.4149 | -103.3691
7.4187 | -103.3269
7.4224 | -103.276
7.4262 | -103.2195
7.4299 | -103.1596
7.4337 | -103.0987
7.4374 | -103.0402
7.4412 | -102.9937
7.4449 | -102.9598
7.4487 | -102.9297
7.4524 | -102.9098
7.4562 | -102.9018
7.4599 | -102.9076
7.4637 | -102.9229
7.4674 | -102.9423
7.4712 | -102.9626
7.4749 | -102.9916
7.4787 | -103.0229
7.4824 | -103.0548
7.4862 | -103.0778
7.4899 | -103.0994
7.4937 | -103.1209
7.4974 | -103.1427
7.5012 | -103.1612
7.5049 | -103.1749
7.5087 | -103.185
7.5124 | -103.1924
7.5162 | -103.2018
7.5199 | -103.2064
7.5237 | -103.2158
7.5274 | -103.2222
7.5312 | -103.2303
7.5349 | -103.2408
7.5387 | -103.2531
7.5424 | -103.2674
7.5462 | -103.2827
7.5499 | -103.2944
7.5537 | -103.31
7.5574 | -103.3339
7.5612 | -103.3531
7.5649 | -103.3697
7.5687 | -103.3879
7.5724 | -103.411
7.5762 | -103.4332
7.5799 | -103.4577
7.5837 | -103.4846
7.5874 | -103.5078
7.5912 | -103.5385
7.5949 | -103.5759
7.5987 | -103.6205
7.6024 | -103.6651
7.6062 | -103.7058
7.6099 | -103.7467
7.6137 | -103.7946
7.6174 | -103.8371
7.6212 | -103.8745
7.6249 | -103.8983
7.6287 | -103.915
7.6324 | -103.933
7.6362 | -103.9366
7.6399 | -103.9298
7.6437 | -103.9092
7.6474 | -103.8758
7.6512 | -103.8371
7.6549 | -103.7941
7.6587 | -103.7508
7.6624 | -103.7099
7.6662 | -103.6704
7.6699 | -103.6379
7.6737 | -103.6161
7.6775 | -103.6011
7.6812 | -103.5961
7.685 | -103.606
7.6887 | -103.6219
7.6925 | -103.6443
7.6962 | -103.6754
7.7 | -103.7062
7.7037 | -103.7354
7.7075 | -103.7573
7.7112 | -103.7733
7.715 | -103.7875
7.7187 | -103.7972
7.7225 | -103.8028
7.7262 | -103.7983
7.73 | -103.7838
7.7337 | -103.7676
7.7375 | -103.7502
7.7412 | -103.736
7.745 | -103.7193
7.7487 | -103.7015
7.7525 | -103.6886
7.7562 | -103.6767
7.76 | -103.673
7.7637 | -103.6742
7.7675 | -103.674
7.7712 | -103.6702
7.775 | -103.6682
7.7787 | -103.6669
7.7825 | -103.6622
7.7862 | -103.655
7.79 | -103.6486
7.7937 | -103.6391
7.7975 | -103.6287
7.8012 | -103.619
7.805 | -103.6152
7.8087 | -103.6179
7.8125 | -103.6204
7.8162 | -103.6319
7.82 | -103.6483
7.8237 | -103.6663
7.8275 | -103.6898
7.8312 | -103.7096
7.835 | -103.7267
7.8387 | -103.7398
7.8425 | -103.7477
7.8462 | -103.7512
7.85 | -103.753
7.8537 | -103.7462
7.8575 | -103.7326
7.8612 | -103.7149
7.865 | -103.6884
7.8687 | -103.6559
7.8725 | -103.6209
7.8762 | -103.5756
7.88 | -103.5279
7.8837 | -103.4857
7.8875 | -103.4482
7.8912 | -103.4056
7.895 | -103.3645
7.8987 | -103.3339
7.9025 | -103.3116
7.9062 | -103.2951
7.91 | -103.2892
7.9137 | -103.2906
7.9175 | -103.301
7.9212 | -103.3227
7.925 | -103.3489
7.9287 | -103.3837
7.9325 | -103.4271
7.9362 | -103.4772
7.94 | -103.5292
7.9437 | -103.582
7.9475 | -103.6399
7.9512 | -103.698
7.955 | -103.7528
7.9587 | -103.7997
7.9625 | -103.8366
7.9662 | -103.8621
7.97 | -103.8866
7.9737 | -103.9072
7.9775 | -103.9121
7.9812 | -103.9121
7.985 | -103.9076
7.9887 | -103.8967
7.9925 | -103.8862
7.9962 | -103.8746
8 | -103.8578
8.0038 | -103.8455
8.0075 | -103.8395
8.0113 | -103.8357
8.015 | -103.8331
8.0188 | -103.831
8.0225 | -103.8355
8.0263 | -103.8395
8.03 | -103.8413
8.0338 | -103.8394
8.0375 | -103.8373
8.0413 | -103.8376
8.045 | -103.8309
8.0488 | -103.8183
8.0525 | -103.8035
8.0563 | -103.7872
8.06 | -103.7678
8.0638 | -103.7488
8.0675 | -103.7309
8.0713 | -103.7155
8.075 | -103.6991
8.0788 | -103.6813
8.0825 | -103.6665
8.0863 | -103.649
8.09 | -103.6358
8.0938 | -103.6319
8.0975 | -103.6317
8.1013 | -103.6337
8.105 | -103.6394
8.1088 | -103.6489
8.1125 | -103.6638
8.1163 | -103.6818
8.12 | -103.6931
8.1238 | -103.7054
8.1275 | -103.7183
8.1313 | -103.7331
8.135 | -103.7445
8.1388 | -103.7533
8.1425 | -103.7616
8.1463 | -103.7726
8.15 | -103.7809
8.1538 | -103.7856
8.1575 | -103.7853
8.1613 | -103.7857
8.165 | -103.7881
8.1688 | -103.793
8.1725 | -103.7967
8.1763 | -103.8009
8.18 | -103.8047
8.1838 | -103.8068
8.1875 | -103.8143
8.1913 | -103.8264
8.195 | -103.8433
8.1988 | -103.8656
8.2025 | -103.8882
8.2063 | -103.914
8.21 | -103.9369
8.2138 | -103.9572
8.2175 | -103.9799
8.2213 | -104.0018
8.225 | -104.0195
8.2288 | -104.0331
8.2325 | -104.0389
8.2363 | -104.0482
8.24 | -104.0618
8.2438 | -104.0715
8.2475 | -104.0744
8.2513 | -104.0747
8.255 | -104.0713
8.2588 | -104.0673
8.2625 | -104.0633
8.2663 | -104.0608
8.27 | -104.0592
8.2738 | -104.0647
8.2775 | -104.071
8.2813 | -104.0812
8.285 | -104.0887
8.2888 | -104.096
8.2925 | -104.1051
8.2963 | -104.1158
8.3 | -104.1309
8.3038 | -104.1458
8.3075 | -104.1526
8.3113 | -104.1446
8.315 | -104.1321
8.3188 | -104.1143
8.3225 | -104.0923
8.3263 | -104.0709
8.3301 | -104.0457
8.3338 | -104.0181
8.3376 | -103.9938
8.3413 | -103.9673
8.3451 | -103.9423
8.3488 | -103.9207
8.3526 | -103.901
8.3563 | -103.8878
8.3601 | -103.8806
8.3638 | -103.8739
8.3676 | -103.8719
8.3713 | -103.8812
8.3751 | -103.893
8.3788 | -103.9084
8.3826 | -103.9284
8.3863 | -103.948
8.3901 | -103.9651
8.3938 | -103.9862
8.3976 | -103.998
8.4013 | -104.0097
8.4051 | -104.0177
8.4088 | -104.0255
8.4126 | -104.0279
8.4163 | -104.0237
8.4201 | -104.0167
8.4238 | -104.0039
8.4276 | -103.9917
8.4313 | -103.9824
8.4351 | -103.9777
8.4388 | -103.9698
8.4426 | -103.9604
8.4463 | -103.9519
8.4501 | -103.9428
8.4538 | -103.9394
8.4576 | -103.9387
8.4613 | -103.9382
8.4651 | -103.9373
8.4688 | -103.9361
8.4726 | -103.9298
8.4763 | -103.9211
8.4801 | -103.909
8.4838 | -103.8896
8.4876 | -103.8653
8.4913 | -103.8396
8.4951 | -103.8099
8.4988 | -103.7838
8.5026 | -103.7616
8.5063 | -103.7379
8.5101 | -103.7196
8.5138 | -103.7023
8.5176 | -103.6904
8.5213 | -103.6947
8.5251 | -103.7069
8.5288 | -103.7248
8.5326 | -103.7487
8.5363 | -103.7834
8.5401 | -103.8255
8.5438 | -103.8728
8.5476 | -103.9296
8.5513 | -103.9878
8.5551 | -104.047
8.5588 | -104.1061
8.5626 | -104.1664
8.5663 | -104.2271
8.5701 | -104.2777
8.5738 | -104.324
8.5776 | -104.3599
8.5813 | -104.3922
8.5851 | -104.419
8.5888 | -104.4353
8.5926 | -104.4446
8.5963 | -104.4474
8.6001 | -104.439
8.6038 | -104.4236
8.6076 | -104.4009
8.6113 | -104.3733
8.6151 | -104.3415
8.6188 | -104.3021
8.6226 | -104.254
8.6263 | -104.2088
8.6301 | -104.167
8.6338 | -104.1287
8.6376 | -104.0947
8.6413 | -104.0594
8.6451 | -104.025
8.6489 | -104.0016
8.6526 | -103.9843
8.6564 | -103.9923
8.6601 | -103.9861
8.6639 | -103.9877
8.6676 | -103.9945
8.6714 | -104.0079
8.6751 | -104.0239
8.6789 | -104.0469
8.6826 | -104.0738
8.6864 | -104.1005
8.6901 | -104.1244
8.6939 | -104.1515
8.6976 | -104.1831
8.7014 | -104.2133
8.7051 | -104.2448
8.7089 | -104.2711
8.7126 | -104.2923
8.7164 | -104.3107
8.7201 | -104.3257
8.7239 | -104.3424
8.7276 | -104.3564
8.7314 | -104.3621
8.7351 | -104.3659
8.7389 | -104.3729
8.7426 | -104.3702
8.7464 | -104.3622
8.7501 | -104.3549
8.7539 | -104.341
8.7576 | -104.3238
8.7614 | -104.295
8.7651 | -104.2647
8.7689 | -104.2397
8.7726 | -104.2126
8.7764 | -104.1794
8.7801 | -104.149
8.7839 | -104.1204
8.7876 | -104.0873
8.7914 | -104.0595
8.7951 | -104.0358
8.7989 | -104.0173
8.8026 | -104.0069
8.8064 | -104.0015
8.8101 | -104.0008
8.8139 | -104.0088
8.8176 | -104.0267
8.8214 | -104.0511
8.8251 | -104.0761
8.8289 | -104.1034
8.8326 | -104.1349
8.8364 | -104.1694
8.8401 | -104.2031
8.8439 | -104.2379
8.8476 | -104.2744
8.8514 | -104.3074
8.8551 | -104.3308
8.8589 | -104.3507
8.8626 | -104.3682
8.8664 | -104.3809
8.8701 | -104.3872
8.8739 | -104.3964
8.8776 | -104.4066
8.8814 | -104.4159
8.8851 | -104.4212
8.8889 | -104.429
8.8926 | -104.4462
8.8964 | -104.4647
8.9001 | -104.4755
8.9039 | -104.4904
8.9076 | -104.5069
8.9114 | -104.5233
8.9151 | -104.534
8.9189 | -104.5448
8.9226 | -104.5532
8.9264 | -104.5528
8.9301 | -104.5384
8.9339 | -104.5129
8.9376 | -104.4825
8.9414 | -104.4439
8.9451 | -104.4019
8.9489 | -104.3558
8.9526 | -104.3116
8.9564 | -104.2633
8.9602 | -104.2136
8.9639 | -104.1661
8.9677 | -104.1201
8.9714 | -104.0801
8.9752 | -104.044
8.9789 | -104.0126
8.9827 | -103.989
8.9864 | -103.9734
8.9902 | -103.9614
8.9939 | -103.9589
8.9977 | -103.9599
9.0014 | -103.9658
9.0052 | -103.974
9.0089 | -103.9842
9.0127 | -103.9947
9.0164 | -104.0106
9.0202 | -104.0323
9.0239 | -104.0522
9.0277 | -104.0692
9.0314 | -104.0913
9.0352 | -104.1129
9.0389 | -104.1364
9.0427 | -104.163
9.0464 | -104.1933
9.0502 | -104.2186
9.0539 | -104.24
9.0577 | -104.2629
9.0614 | -104.2911
9.0652 | -104.3163
9.0689 | -104.3384
9.0727 | -104.3582
9.0764 | -104.3692
9.0802 | -104.3741
9.0839 | -104.3764
9.0877 | -104.3763
9.0914 | -104.3741
9.0952 | -104.3673
9.0989 | -104.3587
9.1027 | -104.3492
9.1064 | -104.3353
9.1102 | -104.3181
9.1139 | -104.2989
9.1177 | -104.2772
9.1214 | -104.2595
9.1252 | -104.2423
9.1289 | -104.2263
9.1327 | -104.2112
9.1364 | -104.2001
9.1402 | -104.1921
9.1439 | -104.1856
9.1477 | -104.1807
9.1514 | -104.1704
9.1552 | -104.1643
9.1589 | -104.1649
9.1627 | -104.1659
9.1664 | -104.168
9.1702 | -104.1705
9.1739 | -104.1725
9.1777 | -104.1727
9.1814 | -104.1779
9.1852 | -104.1857
9.1889 | -104.1938
9.1927 | -104.2009
9.1964 | -104.2123
9.2002 | -104.2235
9.2039 | -104.2335
9.2077 | -104.2463
9.2114 | -104.2596
9.2152 | -104.2736
9.2189 | -104.2886
9.2227 | -104.3041
9.2264 | -104.3201
9.2302 | -104.3315
9.2339 | -104.3411
9.2377 | -104.354
9.2414 | -104.3693
9.2452 | -104.3855
9.2489 | -104.4036
9.2527 | -104.4256
9.2564 | -104.4454
9.2602 | -104.4639
9.2639 | -104.484
9.2677 | -104.5087
9.2714 | -104.5319
9.2752 | -104.5563
9.279 | -104.5829
9.2827 | -104.6081
9.2865 | -104.6396
9.2902 | -104.6667
9.294 | -104.6853
9.2977 | -104.6971
9.3015 | -104.7005
9.3052 | -104.6976
9.309 | -104.6917
9.3127 | -104.6843
9.3165 | -104.6694
9.3202 | -104.6473
9.324 | -104.6244
9.3277 | -104.6025
9.3315 | -104.5758
9.3352 | -104.5488
9.339 | -104.5257
9.3427 | -104.5069
9.3465 | -104.4921
9.3502 | -104.4825
9.354 | -104.4805
9.3577 | -104.4825
9.3615 | -104.4907
9.3652 | -104.5014
9.369 | -104.5215
9.3727 | -104.5446
9.3765 | -104.5698
9.3802 | -104.5991
9.384 | -104.6338
9.3877 | -104.67
9.3915 | -104.7061
9.3952 | -104.7467
9.399 | -104.7899
9.4027 | -104.8332
9.4065 | -104.8695
9.4102 | -104.905
9.414 | -104.938
9.4177 | -104.9652
9.4215 | -104.9847
9.4252 | -104.9965
9.429 | -105.006
9.4327 | -105.0144
9.4365 | -105.02
9.4402 | -105.021
9.444 | -105.0156
9.4477 | -105.0056
9.4515 | -104.9934
9.4552 | -104.9794
9.459 | -104.9632
9.4627 | -104.9486
9.4665 | -104.9318
9.4702 | -104.9217
9.474 | -104.9107
9.4777 | -104.8973
9.4815 | -104.888
9.4852 | -104.8778
9.489 | -104.8662
9.4927 | -104.8549
9.4965 | -104.8506
9.5002 | -104.8521
9.504 | -104.858
9.5077 | -104.8625
9.5115 | -104.8689
9.5152 | -104.8735
9.519 | -104.8873
9.5227 | -104.9009
9.5265 | -104.9142
9.5302 | -104.9295
9.534 | -104.9418
9.5377 | -104.955
9.5415 | -104.9667
9.5452 | -104.9812
9.549 | -104.996
9.5527 | -105.0086
9.5565 | -105.0186
9.5602 | -105.0287
9.564 | -105.0368
9.5677 | -105.0416
9.5715 | -105.0431
9.5752 | -105.0433
9.579 | -105.0464
9.5827 | -105.0425
9.5865 | -105.0348
9.5902 | -105.0237
9.594 | -105.009
9.5977 | -104.9917
9.6015 | -104.9672
9.6053 | -104.9433
9.609 | -104.9127
9.6128 | -104.8752
9.6165 | -104.8356
9.6203 | -104.7977
9.624 | -104.7616
9.6278 | -104.7294
9.6315 | -104.6937
9.6353 | -104.6519
9.639 | -104.6196
9.6428 | -104.5929
9.6465 | -104.575
9.6503 | -104.5636
9.654 | -104.5581
9.6578 | -104.5533
9.6615 | -104.557
9.6653 | -104.5694
9.669 | -104.5894
9.6728 | -104.6172
9.6765 | -104.6452
9.6803 | -104.6776
9.684 | -104.709
9.6878 | -104.7404
9.6915 | -104.7716
9.6953 | -104.8025
9.699 | -104.8279
9.7028 | -104.849
9.7065 | -104.8658
9.7103 | -104.881
9.714 | -104.8935
9.7178 | -104.9028
9.7215 | -104.9118
9.7253 | -104.9158
9.729 | -104.9175
9.7328 | -104.9124
9.7365 | -104.8974
9.7403 | -104.8893
9.744 | -104.8793
9.7478 | -104.8666
9.7515 | -104.8558
9.7553 | -104.8487
9.759 | -104.843
9.7628 | -104.8306
9.7665 | -104.8199
9.7703 | -104.8157
9.774 | -104.8199
9.7778 | -104.8301
9.7815 | -104.8393
9.7853 | -104.8483
9.789 | -104.857
9.7928 | -104.8692
9.7965 | -104.8822
9.8003 | -104.897
9.804 | -104.9087
9.8078 | -104.9194
9.8115 | -104.929
9.8153 | -104.936
9.819 | -104.9454
9.8228 | -104.9526
9.8265 | -104.9621
9.8303 | -104.9719
9.834 | -104.9817
9.8378 | -104.9919
9.8415 | -105.0042
9.8453 | -105.0144
9.849 | -105.023
9.8528 | -105.0257
9.8565 | -105.0234
9.8603 | -105.0195
9.864 | -105.0178
9.8678 | -105.0181
9.8715 | -105.0146
9.8753 | -105.008
9.879 | -104.9954
9.8828 | -104.9828
9.8865 | -104.9644
9.8903 | -104.9401
9.894 | -104.9105
9.8978 | -104.8732
9.9015 | -104.8335
9.9053 | -104.7896
9.909 | -104.7502
9.9128 | -104.7088
9.9165 | -104.6668
9.9203 | -104.6263
9.9241 | -104.5888
9.9278 | -104.5587
9.9316 | -104.5355
9.9353 | -104.5207
9.9391 | -104.5121
9.9428 | -104.5108
9.9466 | -104.5171
9.9503 | -104.5277
9.9541 | -104.5411
9.9578 | -104.5559
9.9616 | -104.574
9.9653 | -104.5913
9.9691 | -104.6116
9.9728 | -104.6321
9.9766 | -104.6492
9.9803 | -104.6638
9.9841 | -104.6761
9.9878 | -104.682
9.9916 | -104.6887
9.9953 | -104.6888
9.9991 | -104.6869
10.0028 | -104.6865
10.0066 | -104.6822
10.0103 | -104.6758
10.0141 | -104.666
10.0178 | -104.661
10.0216 | -104.6635
10.0253 | -104.6702
10.0291 | -104.6756
10.0328 | -104.6847
10.0366 | -104.6984
10.0403 | -104.7165
10.0441 | -104.7406
10.0478 | -104.7623
10.0516 | -104.7818
10.0553 | -104.8049
10.0591 | -104.8288
10.0628 | -104.8507
10.0666 | -104.8721
10.0703 | -104.8923
10.0741 | -104.9064
10.0778 | -104.9152
10.0816 | -104.92
10.0853 | -104.9173
10.0891 | -104.9088
10.0928 | -104.8946
10.0966 | -104.8806
10.1003 | -104.8607
10.1041 | -104.8396
10.1078 | -104.8144
10.1116 | -104.7885
10.1153 | -104.7619
10.1191 | -104.736
10.1228 | -104.7144
10.1266 | -104.6951
10.1303 | -104.6817
10.1341 | -104.6757
10.1378 | -104.6675
10.1416 | -104.6657
10.1453 | -104.6684
10.1491 | -104.6717
10.1528 | -104.6824
10.1566 | -104.6975
10.1603 | -104.7157
10.1641 | -104.7274
10.1678 | -104.7396
10.1716 | -104.7555
10.1753 | -104.7718
10.1791 | -104.7915
10.1828 | -104.8128
10.1866 | -104.8309
10.1903 | -104.8475
10.1941 | -104.8689
10.1978 | -104.8899
10.2016 | -104.9026
10.2053 | -104.9121
10.2091 | -104.9214
10.2128 | -104.9277
10.2166 | -104.9302
10.2203 | -104.9317
10.2241 | -104.9354
10.2278 | -104.9349
10.2316 | -104.9349
10.2353 | -104.932
10.2391 | -104.9232
10.2429 | -104.9132
10.2466 | -104.9047
10.2504 | -104.8924
10.2541 | -104.8775
10.2579 | -104.8647
10.2616 | -104.8502
10.2654 | -104.8342
10.2691 | -104.8189
10.2729 | -104.8082
10.2766 | -104.7975
10.2804 | -104.7981
10.2841 | -104.7943
10.2879 | -104.797
10.2916 | -104.8024
10.2954 | -104.8073
10.2991 | -104.815
10.3029 | -104.8239
10.3066 | -104.8387
10.3104 | -104.8574
10.3141 | -104.8796
10.3179 | -104.899
10.3216 | -104.9155
10.3254 | -104.9317
10.3291 | -104.9453
10.3329 | -104.959
10.3366 | -104.9754
10.3404 | -104.9895
10.3441 | -105.001
10.3479 | -105.0166
10.3516 | -105.0321
10.3554 | -105.0446
10.3591 | -105.0514
10.3629 | -105.0552
10.3666 | -105.06
10.3704 | -105.0668
10.3741 | -105.071
10.3779 | -105.07
10.3816 | -105.0659
10.3854 | -105.0616
10.3891 | -105.0549
10.3929 | -105.0413
10.3966 | -105.0256
10.4004 | -105.0133
10.4041 | -105.0017
10.4079 | -104.9863
10.4116 | -104.9714
10.4154 | -104.9562
10.4191 | -104.9456
10.4229 | -104.9315
10.4266 | -104.9173
10.4304 | -104.9078
10.4341 | -104.8969
10.4379 | -104.8862
10.4416 | -104.8736
10.4454 | -104.8654
10.4491 | -104.8601
10.4529 | -104.8507
10.4566 | -104.8406
10.4604 | -104.8284
10.4641 | -104.8227
10.4679 | -104.8194
10.4716 | -104.8158
10.4754 | -104.8105
10.4791 | -104.8056
10.4829 | -104.7983
10.4866 | -104.7919
10.4904 | -104.7847
10.4941 | -104.7792
10.4979 | -104.7727
10.5016 | -104.7661
10.5054 | -104.759
10.5091 | -104.7546
10.5129 | -104.7512
10.5166 | -104.7477
10.5204 | -104.7418
10.5241 | -104.7351
10.5279 | -104.7306
10.5316 | -104.7306
10.5354 | -104.7342
10.5391 | -104.7372
10.5429 | -104.7446
10.5466 | -104.7556
10.5504 | -104.7673
10.5541 | -104.7779
10.5579 | -104.791
10.5617 | -104.8055
10.5654 | -104.8199
10.5692 | -104.8361
10.5729 | -104.8493
10.5767 | -104.8674
10.5804 | -104.885
10.5842 | -104.8991
10.5879 | -104.9119
10.5917 | -104.9248
10.5954 | -104.9361
10.5992 | -104.9367
10.6029 | -104.9362
10.6067 | -104.9396
10.6104 | -104.9431
10.6142 | -104.9434
10.6179 | -104.9386
10.6217 | -104.9361
10.6254 | -104.9299
10.6292 | -104.9244
10.6329 | -104.9178
10.6367 | -104.9129
10.6404 | -104.9054
10.6442 | -104.8994
10.6479 | -104.8916
10.6517 | -104.8849
10.6554 | -104.878
10.6592 | -104.8755
10.6629 | -104.8755
10.6667 | -104.8748
10.6704 | -104.8756
10.6742 | -104.8827
10.6779 | -104.8919
10.6817 | -104.9026
10.6854 | -104.9169
10.6892 | -104.9323
10.6929 | -104.952
10.6967 | -104.9704
10.7004 | -104.9887
10.7042 | -105.0051
10.7079 | -105.0223
10.7117 | -105.0409
10.7154 | -105.0557
10.7192 | -105.068
10.7229 | -105.0801
10.7267 | -105.0929
10.7304 | -105.1004
10.7342 | -105.1066
10.7379 | -105.1119
10.7417 | -105.113
10.7454 | -105.1126
10.7492 | -105.1157
10.7529 | -105.1177
10.7567 | -105.1151
10.7604 | -105.1093
10.7642 | -105.0995
10.7679 | -105.0891
10.7717 | -105.0775
10.7754 | -105.0652
10.7792 | -105.0578
10.7829 | -105.0482
10.7867 | -105.037
10.7904 | -105.0239
10.7942 | -105.0085
10.7979 | -104.9943
10.8017 | -104.9843
10.8054 | -104.9728
10.8092 | -104.9579
10.8129 | -104.9433
10.8167 | -104.9295
10.8204 | -104.9224
10.8242 | -104.9099
10.8279 | -104.8983
10.8317 | -104.8875
10.8354 | -104.8728
10.8392 | -104.8642
10.8429 | -104.8561
10.8467 | -104.8468
10.8504 | -104.8393
10.8542 | -104.8333
10.8579 | -104.8277
10.8617 | -104.8232
10.8654 | -104.8223
10.8692 | -104.8208
10.8729 | -104.8227
10.8767 | -104.8283
10.8804 | -104.8351
10.8842 | -104.8397
10.888 | -104.8485
10.8917 | -104.8613
10.8955 | -104.8733
10.8992 | -104.8872
10.903 | -104.9042
10.9067 | -104.9236
10.9105 | -104.9473
10.9142 | -104.9694
10.918 | -104.9933
10.9217 | -105.0221
10.9255 | -105.0507
10.9292 | -105.0772
10.933 | -105.1051
10.9367 | -105.1342
10.9405 | -105.1671
10.9442 | -105.2012
10.948 | -105.2331
10.9517 | -105.2675
10.9555 | -105.3012
10.9592 | -105.3299
10.963 | -105.3601
10.9667 | -105.392
10.9705 | -105.4155
10.9742 | -105.4326
10.978 | -105.4414
10.9817 | -105.452
10.9855 | -105.463
10.9892 | -105.4678
10.993 | -105.4659
10.9967 | -105.4635
11.0005 | -105.4531
11.0042 | -105.4402
11.008 | -105.4284
11.0117 | -105.409
11.0155 | -105.3875
11.0192 | -105.3642
11.023 | -105.3415
11.0267 | -105.3128
11.0305 | -105.2798
11.0342 | -105.2455
11.038 | -105.2119
11.0417 | -105.1787
11.0455 | -105.1487
11.0492 | -105.1204
11.053 | -105.0901
11.0567 | -105.0641
11.0605 | -105.0412
11.0642 | -105.0158
11.068 | -104.9926
11.0717 | -104.9688
11.0755 | -104.9495
11.0792 | -104.9356
11.083 | -104.9199
11.0867 | -104.9102
11.0905 | -104.9048
11.0942 | -104.8997
11.098 | -104.8919
11.1017 | -104.8858
11.1055 | -104.8845
11.1092 | -104.8859
11.113 | -104.8861
11.1167 | -104.8881
11.1205 | -104.8932
11.1242 | -104.901
11.128 | -104.9118
11.1317 | -104.9261
11.1355 | -104.9428
11.1392 | -104.9631
11.143 | -104.9832
11.1467 | -105.0013
11.1505 | -105.0219
11.1542 | -105.0423
11.158 | -105.0655
11.1617 | -105.0887
11.1655 | -105.1112
11.1692 | -105.1315
11.173 | -105.155
11.1767 | -105.1725
11.1805 | -105.1863
11.1842 | -105.1968
11.188 | -105.2034
11.1917 | -105.2144
11.1955 | -105.2234
11.1992 | -105.2228
11.203 | -105.2197
11.2068 | -105.2125
11.2105 | -105.1989
11.2143 | -105.1843
11.218 | -105.1699
11.2218 | -105.1532
11.2255 | -105.138
11.2293 | -105.1177
11.233 | -105.1015
11.2368 | -105.0855
11.2405 | -105.0646
11.2443 | -105.0474
11.248 | -105.0308
11.2518 | -105.0181
11.2555 | -105.0102
11.2593 | -105.0034
11.263 | -104.9979
11.2668 | -104.9949
11.2705 | -104.9976
11.2743 | -105.0046
11.278 | -105.0135
11.2818 | -105.0274
11.2855 | -105.0415
11.2893 | -105.059
11.293 | -105.0807
11.2968 | -105.1039
11.3005 | -105.1249
11.3043 | -105.1445
11.308 | -105.1674
11.3118 | -105.1888
11.3155 | -105.2116
11.3193 | -105.2329
11.323 | -105.2508
11.3268 | -105.2664
11.3305 | -105.2825
11.3343 | -105.296
11.338 | -105.3093
11.3418 | -105.322
11.3455 | -105.3299
11.3493 | -105.3337
11.353 | -105.3335
11.3568 | -105.3362
11.3605 | -105.3353
11.3643 | -105.3398
11.368 | -105.3428
11.3718 | -105.3457
11.3755 | -105.3488
11.3793 | -105.3511
11.383 | -105.3505
11.3868 | -105.3497
11.3905 | -105.3483
11.3943 | -105.3487
11.398 | -105.3471
11.4018 | -105.3466
11.4055 | -105.3476
11.4093 | -105.3499
11.413 | -105.3514
11.4168 | -105.3536
11.4205 | -105.3532
11.4243 | -105.3585
11.428 | -105.3666
11.4318 | -105.3743
11.4355 | -105.3791
11.4393 | -105.3868
11.443 | -105.3936
11.4468 | -105.3971
11.4505 | -105.4042
11.4543 | -105.4105
11.458 | -105.4159
11.4618 | -105.4191
11.4655 | -105.4256
11.4693 | -105.4343
11.473 | -105.4395
11.4768 | -105.4414
11.4805 | -105.4363
11.4843 | -105.4295
11.488 | -105.4214
11.4918 | -105.4065
11.4955 | -105.3886
11.4993 | -105.3753
11.503 | -105.3617
11.5068 | -105.3463
11.5105 | -105.3327
11.5143 | -105.3224
11.5181 | -105.3131
11.5218 | -105.3029
11.5256 | -105.2952
11.5293 | -105.2877
11.5331 | -105.2839
11.5368 | -105.281
11.5406 | -105.2819
11.5443 | -105.2834
11.5481 | -105.2852
11.5518 | -105.2878
11.5556 | -105.2931
11.5593 | -105.299
11.5631 | -105.3068
11.5668 | -105.3135
11.5706 | -105.3183
11.5743 | -105.3277
11.5781 | -105.3366
11.5818 | -105.3472
11.5856 | -105.3563
11.5893 | -105.3668
11.5931 | -105.3771
11.5968 | -105.385
11.6006 | -105.3935
11.6043 | -105.4013
11.6081 | -105.4028
11.6118 | -105.4056
11.6156 | -105.4098
11.6193 | -105.4073
11.6231 | -105.4008
11.6268 | -105.395
11.6306 | -105.3913
11.6343 | -105.3817
11.6381 | -105.3662
11.6418 | -105.3496
11.6456 | -105.3356
11.6493 | -105.3206
11.6531 | -105.3069
11.6568 | -105.2979
11.6606 | -105.2896
11.6643 | -105.2757
11.6681 | -105.2574
11.6718 | -105.2404
11.6756 | -105.2232
11.6793 | -105.2061
11.6831 | -105.1894
11.6868 | -105.1714
11.6906 | -105.1569
11.6943 | -105.1465
11.6981 | -105.1376
11.7018 | -105.1284
11.7056 | -105.1245
11.7093 | -105.1255
11.7131 | -105.1325
11.7168 | -105.1409
11.7206 | -105.1543
11.7243 | -105.1718
11.7281 | -105.1911
11.7318 | -105.2124
11.7356 | -105.2395
11.7393 | -105.2696
11.7431 | -105.3007
11.7468 | -105.3368
11.7506 | -105.3764
11.7543 | -105.4154
11.7581 | -105.4555
11.7618 | -105.4941
11.7656 | -105.529
11.7693 | -105.5636
11.7731 | -105.6003
11.7768 | -105.6368
11.7806 | -105.6643
11.7843 | -105.6841
11.7881 | -105.7001
11.7918 | -105.7136
11.7956 | -105.7188
11.7993 | -105.7156
11.8031 | -105.7059
11.8068 | -105.6952
11.8106 | -105.6782
11.8143 | -105.651
11.8181 | -105.6217
11.8218 | -105.5851
11.8256 | -105.5426
11.8293 | -105.4988
11.8331 | -105.4551
11.8369 | -105.4102
11.8406 | -105.3679
11.8444 | -105.324
11.8481 | -105.2805
11.8519 | -105.243
11.8556 | -105.2058
11.8594 | -105.1691
11.8631 | -105.1325
11.8669 | -105.1015
11.8706 | -105.0722
11.8744 | -105.0467
11.8781 | -105.0259
11.8819 | -105.0086
11.8856 | -104.9983
11.8894 | -104.991
11.8931 | -104.9855
11.8969 | -104.9862
11.9006 | -105.0046
11.9044 | -105.0149
11.9081 | -105.027
11.9119 | -105.0429
11.9156 | -105.0626
11.9194 | -105.0837
11.9231 | -105.1052
11.9269 | -105.1334
11.9306 | -105.1677
11.9344 | -105.2053
11.9381 | -105.246
11.9419 | -105.2842
11.9456 | -105.3234
11.9494 | -105.3622
11.9531 | -105.4036
11.9569 | -105.4457
11.9606 | -105.486
11.9644 | -105.5225
11.9681 | -105.5548
11.9719 | -105.5876
11.9756 | -105.6178
11.9794 | -105.64
11.9831 | -105.6611
11.9869 | -105.6798
11.9906 | -105.6918
11.9944 | -105.6998
11.9981 | -105.7025
12.0019 | -105.7015
12.0056 | -105.6962
12.0094 | -105.6854
12.0131 | -105.6703
12.0169 | -105.655
12.0206 | -105.6349
12.0244 | -105.6141
12.0281 | -105.5954
12.0319 | -105.5784
12.0356 | -105.562
12.0394 | -105.5432
12.0431 | -105.5255
12.0469 | -105.508
12.0506 | -105.4926
12.0544 | -105.4828
12.0581 | -105.4741
12.0619 | -105.4678
12.0656 | -105.4679
12.0694 | -105.4691
12.0731 | -105.47
12.0769 | -105.473
12.0806 | -105.4757
12.0844 | -105.4744
12.0881 | -105.4741
12.0919 | -105.4779
12.0956 | -105.4836
12.0994 | -105.4897
12.1031 | -105.4924
12.1069 | -105.4895
12.1106 | -105.4849
12.1144 | -105.4805
12.1181 | -105.4767
12.1219 | -105.4696
12.1256 | -105.4626
12.1294 | -105.456
12.1331 | -105.4494
12.1369 | -105.4428
12.1406 | -105.435
12.1444 | -105.4248
12.1481 | -105.4106
12.1519 | -105.4034
12.1556 | -105.397
12.1594 | -105.3868
12.1631 | -105.378
12.1669 | -105.371
12.1707 | -105.366
12.1744 | -105.3631
12.1782 | -105.3587
12.1819 | -105.3515
12.1857 | -105.3438
12.1894 | -105.3346
12.1932 | -105.3248
12.1969 | -105.3169
12.2007 | -105.3046
12.2044 | -105.292
12.2082 | -105.2789
12.2119 | -105.2673
12.2157 | -105.2544
12.2194 | -105.2444
12.2232 | -105.2333
12.2269 | -105.2166
12.2307 | -105.2004
12.2344 | -105.1866
12.2382 | -105.1749
12.2419 | -105.1649
12.2457 | -105.1614
12.2494 | -105.1621
12.2532 | -105.1639
12.2569 | -105.1701
12.2607 | -105.1808
12.2644 | -105.1919
12.2682 | -105.2037
12.2719 | -105.2166
12.2757 | -105.232
12.2794 | -105.2523
12.2832 | -105.2751
12.2869 | -105.3009
12.2907 | -105.3259
12.2944 | -105.3506
12.2982 | -105.376
12.3019 | -105.4034
12.3057 | -105.4272
12.3094 | -105.449
12.3132 | -105.4691
12.3169 | -105.4901
12.3207 | -105.5125
12.3244 | -105.5356
12.3282 | -105.5512
12.3319 | -105.5676
12.3357 | -105.5829
12.3394 | -105.5934
12.3432 | -105.6035
12.3469 | -105.6131
12.3507 | -105.6197
12.3544 | -105.6266
12.3582 | -105.6323
12.3619 | -105.6353
12.3657 | -105.6359
12.3694 | -105.6353
12.3732 | -105.6323
12.3769 | -105.6277
12.3807 | -105.6223
12.3844 | -105.6169
12.3882 | -105.6101
12.3919 | -105.6074
12.3957 | -105.6006
12.3994 | -105.5909
12.4032 | -105.5843
12.4069 | -105.5772
12.4107 | -105.5693
12.4144 | -105.5631
12.4182 | -105.5556
12.4219 | -105.5469
12.4257 | -105.541
12.4294 | -105.5361
12.4332 | -105.5323
12.4369 | -105.5327
12.4407 | -105.5304
12.4444 | -105.5263
12.4482 | -105.5196
12.4519 | -105.5201
12.4557 | -105.5225
12.4594 | -105.5229
12.4632 | -105.5206
12.4669 | -105.519
12.4707 | -105.5205
12.4744 | -105.5238
12.4782 | -105.5264
12.4819 | -105.527
12.4857 | -105.5298
12.4895 | -105.533
12.4932 | -105.5353
12.497 | -105.5363
12.5007 | -105.5367
12.5045 | -105.5364
12.5082 | -105.536
12.512 | -105.5345
12.5157 | -105.5335
12.5195 | -105.5333
12.5232 | -105.5296
12.527 | -105.523
12.5307 | -105.5228
12.5345 | -105.5218
12.5382 | -105.5226
12.542 | -105.522
12.5457 | -105.5207
12.5495 | -105.5177
12.5532 | -105.5142
12.557 | -105.5098
12.5607 | -105.5046
12.5645 | -105.4974
12.5682 | -105.4908
12.572 | -105.4864
12.5757 | -105.4796
12.5795 | -105.4736
12.5832 | -105.4691
12.587 | -105.4602
12.5907 | -105.4507
12.5945 | -105.4423
12.5982 | -105.4346
12.602 | -105.4289
12.6057 | -105.4203
12.6095 | -105.4091
12.6132 | -105.4001
12.617 | -105.3922
12.6207 | -105.3856
12.6245 | -105.3812
12.6282 | -105.3762
12.632 | -105.3695
12.6357 | -105.3607
12.6395 | -105.3523
12.6432 | -105.3431
12.647 | -105.3377
12.6507 | -105.3339
12.6545 | -105.3278
12.6582 | -105.3223
12.662 | -105.3178
12.6657 | -105.314
12.6695 | -105.3127
12.6732 | -105.3106
12.677 | -105.3085
12.6807 | -105.308
12.6845 | -105.3128
12.6882 | -105.3152
12.692 | -105.3185
12.6957 | -105.3236
12.6995 | -105.3317
12.7032 | -105.3409
12.707 | -105.3506
12.7107 | -105.3611
12.7145 | -105.3679
12.7182 | -105.376
12.722 | -105.3851
12.7257 | -105.3963
12.7295 | -105.4061
12.7332 | -105.4134
12.737 | -105.4211
12.7407 | -105.4276
12.7445 | -105.4331
12.7482 | -105.4392
12.752 | -105.4454
12.7557 | -105.4471
12.7595 | -105.4523
12.7632 | -105.4552
12.767 | -105.4548
12.7707 | -105.4573
12.7745 | -105.4571
12.7782 | -105.454
12.782 | -105.4525
12.7857 | -105.4484
12.7895 | -105.4468
12.7932 | -105.4479
12.797 | -105.4496
12.8008 | -105.4509
12.8045 | -105.4541
12.8083 | -105.46
12.812 | -105.467
12.8158 | -105.4741
12.8195 | -105.4829
12.8233 | -105.4937
12.827 | -105.5079
12.8308 | -105.5224
12.8345 | -105.5332
12.8383 | -105.549
12.842 | -105.565
12.8458 | -105.582
12.8495 | -105.5995
12.8533 | -105.6161
12.857 | -105.637
12.8608 | -105.6567
12.8645 | -105.6718
12.8683 | -105.6843
12.872 | -105.6998
12.8758 | -105.7146
12.8795 | -105.7269
12.8833 | -105.7352
12.887 | -105.744
12.8908 | -105.7541
12.8945 | -105.7606
12.8983 | -105.765
12.902 | -105.7656
12.9058 | -105.7651
12.9095 | -105.764
12.9133 | -105.7646
12.917 | -105.7655
12.9208 | -105.7611
12.9245 | -105.7587
12.9283 | -105.7575
12.932 | -105.7538
12.9358 | -105.7507
12.9395 | -105.7441
12.9433 | -105.7364
12.947 | -105.734
12.9508 | -105.7326
12.9545 | -105.7296
12.9583 | -105.727
12.962 | -105.7234
12.9658 | -105.719
12.9695 | -105.7183
12.9733 | -105.7144
12.977 | -105.7104
12.9808 | -105.7076
12.9845 | -105.7172
12.9883 | -105.7187
12.992 | -105.7223
12.9958 | -105.7252
12.9995 | -105.7306
13.0033 | -105.7397
13.007 | -105.7479
13.0108 | -105.7565
13.0145 | -105.7622
13.0183 | -105.7689
13.022 | -105.7767
13.0258 | -105.7825
13.0295 | -105.7848
13.0333 | -105.7868
13.037 | -105.7899
13.0408 | -105.7903
13.0445 | -105.7892
13.0483 | -105.7879
13.052 | -105.7855
13.0558 | -105.7795
13.0595 | -105.7717
13.0633 | -105.7614
13.067 | -105.7517
13.0708 | -105.7442
13.0745 | -105.7322
13.0783 | -105.7233
13.082 | -105.7168
13.0858 | -105.7085
13.0895 | -105.7027
13.0933 | -105.6955
13.097 | -105.6887
13.1008 | -105.6807
13.1045 | -105.6754
13.1083 | -105.6759
13.112 | -105.6768
13.1158 | -105.6796
13.1196 | -105.6854
13.1233 | -105.692
13.1271 | -105.7043
13.1308 | -105.7167
13.1346 | -105.7295
13.1383 | -105.745
13.1421 | -105.7634
13.1458 | -105.7861
13.1496 | -105.8114
13.1533 | -105.8373
13.1571 | -105.8606
13.1608 | -105.8836
13.1646 | -105.9055
13.1683 | -105.9299
13.1721 | -105.9488
13.1758 | -105.9672
13.1796 | -105.9845
13.1833 | -105.9966
13.1871 | -106.0043
13.1908 | -106.0111
13.1946 | -106.0151
13.1983 | -106.0148
13.2021 | -106.0082
13.2058 | -105.9978
13.2096 | -105.9832
13.2133 | -105.9626
13.2171 | -105.9361
13.2208 | -105.9132
13.2246 | -105.8866
13.2283 | -105.8533
13.2321 | -105.82
13.2358 | -105.7821
13.2396 | -105.741
13.2433 | -105.6999
13.2471 | -105.6572
13.2508 | -105.6123
13.2546 | -105.5677
13.2583 | -105.5211
13.2621 | -105.4733
13.2658 | -105.4274
13.2696 | -105.3792
13.2733 | -105.3309
13.2771 | -105.2866
13.2808 | -105.2421
13.2846 | -105.1985
13.2883 | -105.1567
13.2921 | -105.1165
13.2958 | -105.0793
13.2996 | -105.0417
13.3033 | -105.0055
13.3071 | -104.9729
13.3108 | -104.9439
13.3146 | -104.9168
13.3183 | -104.8891
13.3221 | -104.8612
13.3258 | -104.8346
13.3296 | -104.8106
13.3333 | -104.7893
13.3371 | -104.7666
13.3408 | -104.7445
13.3446 | -104.7246
13.3483 | -104.706
13.3521 | -104.6901
13.3558 | -104.6717
13.3596 | -104.6454
13.3633 | -104.6188
13.3671 | -104.5987
13.3708 | -104.5801
13.3746 | -104.56
13.3783 | -104.5419
13.3821 | -104.5196
13.3858 | -104.4959
13.3896 | -104.4748
13.3933 | -104.4546
13.3971 | -104.4353
13.4008 | -104.4149
13.4046 | -104.3944
13.4083 | -104.3767
13.4121 | -104.3573
13.4158 | -104.34
13.4196 | -104.3275
13.4233 | -104.3158
13.4271 | -104.3053
13.4308 | -104.3011
13.4346 | -104.3005
13.4383 | -104.3033
13.4421 | -104.3111
13.4459 | -104.3234
13.4496 | -104.3415
13.4534 | -104.3637
13.4571 | -104.3901
13.4609 | -104.4211
13.4646 | -104.4572
13.4684 | -104.4986
13.4721 | -104.5455
13.4759 | -104.5948
13.4796 | -104.6454
13.4834 | -104.7004
13.4871 | -104.7547
13.4909 | -104.8109
13.4946 | -104.8677
13.4984 | -104.9226
13.5021 | -104.9784
13.5059 | -105.0328
13.5096 | -105.0866
13.5134 | -105.1395
13.5171 | -105.1855
13.5209 | -105.2237
13.5246 | -105.2379
13.5284 | -105.2702
13.5321 | -105.2968
13.5359 | -105.3152
13.5396 | -105.3291
13.5434 | -105.3409
13.5471 | -105.3481
13.5509 | -105.3523
13.5546 | -105.3532
13.5584 | -105.3489
13.5621 | -105.3369
13.5659 | -105.3212
13.5696 | -105.3052
13.5734 | -105.2871
13.5771 | -105.2681
13.5809 | -105.2511
13.5846 | -105.2306
13.5884 | -105.2125
13.5921 | -105.1958
13.5959 | -105.1801
13.5996 | -105.1657
13.6034 | -105.1511
13.6071 | -105.1388
13.6109 | -105.1301
13.6146 | -105.1238
13.6184 | -105.122
13.6221 | -105.1249
13.6259 | -105.1286
13.6296 | -105.1355
13.6334 | -105.1421
13.6371 | -105.1492
13.6409 | -105.1599
13.6446 | -105.171
13.6484 | -105.1854
13.6521 | -105.201
13.6559 | -105.2168
13.6596 | -105.2339
13.6634 | -105.2495
13.6671 | -105.2667
13.6709 | -105.2863
13.6746 | -105.3057
13.6784 | -105.3223
13.6821 | -105.339
13.6859 | -105.3533
13.6896 | -105.3647
13.6934 | -105.3756
13.6971 | -105.3852
13.7009 | -105.394
13.7046 | -105.4009
13.7084 | -105.4074
13.7121 | -105.4125
13.7159 | -105.4192
13.7196 | -105.4223
13.7234 | -105.4192
13.7271 | -105.4178
13.7309 | -105.4228
13.7346 | -105.4286
13.7384 | -105.434
13.7421 | -105.4364
13.7459 | -105.4396
13.7496 | -105.4463
13.7534 | -105.456
13.7571 | -105.4681
13.7609 | -105.4784
13.7647 | -105.4911
13.7684 | -105.506
13.7722 | -105.5208
13.7759 | -105.5367
13.7797 | -105.5523
13.7834 | -105.5705
13.7872 | -105.5878
13.7909 | -105.605
13.7947 | -105.6219
13.7984 | -105.6393
13.8022 | -105.6599
13.8059 | -105.6773
13.8097 | -105.6908
13.8134 | -105.7056
13.8172 | -105.7208
13.8209 | -105.7355
13.8247 | -105.7484
13.8284 | -105.7598
13.8322 | -105.7704
13.8359 | -105.782
13.8397 | -105.7908
13.8434 | -105.7983
13.8472 | -105.8027
13.8509 | -105.8067
13.8547 | -105.8092
13.8584 | -105.8124
13.8622 | -105.8154
13.8659 | -105.8117
13.8697 | -105.8048
13.8734 | -105.796
13.8772 | -105.7892
13.8809 | -105.784
13.8847 | -105.7784
13.8884 | -105.7715
13.8922 | -105.7626
13.8959 | -105.7494
13.8997 | -105.7351
13.9034 | -105.7211
13.9072 | -105.7071
13.9109 | -105.6924
13.9147 | -105.6767
13.9184 | -105.6607
13.9222 | -105.646
13.9259 | -105.6272
13.9297 | -105.6085
13.9334 | -105.5919
13.9372 | -105.5738
13.9409 | -105.5565
13.9447 | -105.5387
13.9484 | -105.5196
13.9522 | -105.4977
13.9559 | -105.4777
13.9597 | -105.4538
13.9634 | -105.4243
13.9672 | -105.3911
13.9709 | -105.3639
13.9747 | -105.3393
13.9784 | -105.3149
13.9822 | -105.2923
13.9859 | -105.2709
13.9897 | -105.2495
13.9934 | -105.227
13.9972 | -105.2058
14.0009 | -105.182
14.0047 | -105.1585
14.0084 | -105.1389
14.0122 | -105.1168
14.0159 | -105.0933
14.0197 | -105.0726
14.0234 | -105.0546
14.0272 | -105.0385
14.0309 | -105.0225
14.0347 | -105.0075
14.0384 | -104.9947
14.0422 | -104.984
14.0459 | -104.9753
14.0497 | -104.9702
14.0534 | -104.9668
14.0572 | -104.964
14.0609 | -104.9632
14.0647 | -104.9777
14.0684 | -104.9836
14.0722 | -104.9906
14.0759 | -105.0007
14.0797 | -105.0139
14.0835 | -105.026
14.0872 | -105.0406
14.091 | -105.0563
14.0947 | -105.0723
14.0985 | -105.0897
14.1022 | -105.1073
14.106 | -105.123
14.1097 | -105.138
14.1135 | -105.1543
14.1172 | -105.1683
14.121 | -105.1821
14.1247 | -105.1973
14.1285 | -105.2135
14.1322 | -105.2257
14.136 | -105.2365
14.1397 | -105.2487
14.1435 | -105.262
14.1472 | -105.2769
14.151 | -105.2902
14.1547 | -105.306
14.1585 | -105.3196
14.1622 | -105.3351
14.166 | -105.3554
14.1697 | -105.379
14.1735 | -105.3961
14.1772 | -105.4096
14.181 | -105.4242
14.1847 | -105.4395
14.1885 | -105.4551
14.1922 | -105.4729
14.196 | -105.494
14.1997 | -105.5159
14.2035 | -105.5351
14.2072 | -105.5556
14.211 | -105.5739
14.2147 | -105.5947
14.2185 | -105.6158
14.2222 | -105.6332
14.226 | -105.6484
14.2297 | -105.6632
14.2335 | -105.6821
14.2372 | -105.6979
14.241 | -105.7111
14.2447 | -105.7224
14.2485 | -105.7338
14.2522 | -105.7436
14.256 | -105.7549
14.2597 | -105.7654
14.2635 | -105.775
14.2672 | -105.7805
14.271 | -105.785
14.2747 | -105.7908
14.2785 | -105.7948
14.2822 | -105.7959
14.286 | -105.7927
14.2897 | -105.7873
14.2935 | -105.7853
14.2972 | -105.7843
14.301 | -105.7829
14.3047 | -105.7815
14.3085 | -105.7773
14.3122 | -105.7719
14.316 | -105.7617
14.3197 | -105.7535
14.3235 | -105.7462
14.3272 | -105.7396
14.331 | -105.731
14.3347 | -105.721
14.3385 | -105.7122
14.3422 | -105.7051
14.346 | -105.6952
14.3497 | -105.6853
14.3535 | -105.6771
14.3572 | -105.6702
14.361 | -105.6645
14.3647 | -105.6604
14.3685 | -105.6587
14.3722 | -105.6552
14.376 | -105.65
14.3797 | -105.6464
14.3835 | -105.6463
14.3872 | -105.645
14.391 | -105.6437
14.3947 | -105.6444
14.3985 | -105.6449
14.4023 | -105.6461
14.406 | -105.6489
14.4098 | -105.6554
14.4135 | -105.6625
14.4173 | -105.6686
14.421 | -105.6752
14.4248 | -105.6826
14.4285 | -105.6918
14.4323 | -105.7029
14.436 | -105.7112
14.4398 | -105.7141
14.4435 | -105.7168
14.4473 | -105.7225
14.451 | -105.7305
14.4548 | -105.7366
14.4585 | -105.7418
14.4623 | -105.7451
14.466 | -105.7477
14.4698 | -105.7497
14.4735 | -105.7519
14.4773 | -105.7501
14.481 | -105.7441
14.4848 | -105.7378
14.4885 | -105.7339
14.4923 | -105.7277
14.496 | -105.7187
14.4998 | -105.7088
14.5035 | -105.7008
14.5073 | -105.6936
14.511 | -105.683
14.5148 | -105.6704
14.5185 | -105.6609
14.5223 | -105.6555
14.526 | -105.6456
14.5298 | -105.6351
14.5335 | -105.6245
14.5373 | -105.6151
14.541 | -105.6066
14.5448 | -105.5985
14.5485 | -105.5911
14.5523 | -105.5863
14.556 | -105.582
14.5598 | -105.577
14.5635 | -105.5728
14.5673 | -105.5676
14.571 | -105.5641
14.5748 | -105.561
14.5785 | -105.5603
14.5823 | -105.562
14.586 | -105.5663
14.5898 | -105.5685
14.5935 | -105.571
14.5973 | -105.5714
14.601 | -105.5734
14.6048 | -105.5793
14.6085 | -105.5797
14.6123 | -105.5816
14.616 | -105.5815
14.6198 | -105.5789
14.6235 | -105.5748
14.6273 | -105.5715
14.631 | -105.5662
14.6348 | -105.5597
14.6385 | -105.5509
14.6423 | -105.5427
14.646 | -105.5362
14.6498 | -105.5313
14.6535 | -105.5254
14.6573 | -105.5161
14.661 | -105.5041
14.6648 | -105.4943
14.6685 | -105.4837
14.6723 | -105.4737
14.676 | -105.4658
14.6798 | -105.459
14.6835 | -105.4535
14.6873 | -105.4513
14.691 | -105.4509
14.6948 | -105.4524
14.6985 | -105.4528
14.7023 | -105.4572
14.706 | -105.468
14.7098 | -105.4747
14.7135 | -105.4834
14.7173 | -105.4946
14.721 | -105.5119
14.7248 | -105.5335
14.7286 | -105.5532
14.7323 | -105.5703
14.7361 | -105.5897
14.7398 | -105.6126
14.7436 | -105.6351
14.7473 | -105.6576
14.7511 | -105.6795
14.7548 | -105.6985
14.7586 | -105.7161
14.7623 | -105.734
14.7661 | -105.7512
14.7698 | -105.7667
14.7736 | -105.7792
14.7773 | -105.7901
14.7811 | -105.7979
14.7848 | -105.8024
14.7886 | -105.8041
14.7923 | -105.8016
14.7961 | -105.7952
14.7998 | -105.7899
14.8036 | -105.7826
14.8073 | -105.7777
14.8111 | -105.7706
14.8148 | -105.7583
14.8186 | -105.7449
14.8223 | -105.7316
14.8261 | -105.7159
14.8298 | -105.7032
14.8336 | -105.6922
14.8373 | -105.6813
14.8411 | -105.6664
14.8448 | -105.6486
14.8486 | -105.6322
14.8523 | -105.6183
14.8561 | -105.6002
14.8598 | -105.5793
14.8636 | -105.5611
14.8673 | -105.5445
14.8711 | -105.5269
14.8748 | -105.5115
14.8786 | -105.4945
14.8823 | -105.4796
14.8861 | -105.4634
14.8898 | -105.4462
14.8936 | -105.4316
14.8973 | -105.4192
14.9011 | -105.4076
14.9048 | -105.3951
14.9086 | -105.3823
14.9123 | -105.3718
14.9161 | -105.3612
14.9198 | -105.3521
14.9236 | -105.3465
14.9273 | -105.3397
14.9311 | -105.3362
14.9348 | -105.3327
14.9386 | -105.3294
14.9423 | -105.3291
14.9461 | -105.3318
14.9498 | -105.3321
14.9536 | -105.3288
14.9573 | -105.3304
14.9611 | -105.3314
14.9648 | -105.3326
14.9686 | -105.3378
14.9723 | -105.3418
14.9761 | -105.3418
14.9798 | -105.337
14.9836 | -105.3349
14.9873 | -105.3322
14.9911 | -105.3299
14.9948 | -105.3275
14.9986 | -105.3232
15.0023 | -105.3139
15.0061 | -105.3039
15.0098 | -105.2905
15.0136 | -105.278
15.0173 | -105.2629
15.0211 | -105.2428
15.0248 | -105.2196
15.0286 | -105.1971
15.0323 | -105.1748
15.0361 | -105.1493
15.0398 | -105.1219
15.0436 | -105.0901
15.0474 | -105.0574
15.0511 | -105.027
15.0549 | -104.9969
15.0586 | -104.969
15.0624 | -104.9416
15.0661 | -104.9138
15.0699 | -104.8886
15.0736 | -104.8684
15.0774 | -104.8495
15.0811 | -104.8317
15.0849 | -104.8189
15.0886 | -104.8064
15.0924 | -104.7976
15.0961 | -104.793
15.0999 | -104.7896
15.1036 | -104.7878
15.1074 | -104.7876
15.1111 | -104.7906
15.1149 | -104.796
15.1186 | -104.8073
15.1224 | -104.8207
15.1261 | -104.8348
15.1299 | -104.8528
15.1336 | -104.8707
15.1374 | -104.8913
15.1411 | -104.9163
15.1449 | -104.9424
15.1486 | -104.9729
15.1524 | -105
15.1561 | -105.0293
15.1599 | -105.0565
15.1636 | -105.0835
15.1674 | -105.1127
15.1711 | -105.1446
15.1749 | -105.1713
15.1786 | -105.1964
15.1824 | -105.2202
15.1861 | -105.2426
15.1899 | -105.2647
15.1936 | -105.2863
15.1974 | -105.3085
15.2011 | -105.3267
15.2049 | -105.3446
15.2086 | -105.3563
15.2124 | -105.3641
15.2161 | -105.3705
15.2199 | -105.3791
15.2236 | -105.3859
15.2274 | -105.3916
15.2311 | -105.3951
15.2349 | -105.3987
15.2386 | -105.4022
15.2424 | -105.4055
15.2461 | -105.409
15.2499 | -105.4096
15.2536 | -105.409
15.2574 | -105.4101
15.2611 | -105.4122
15.2649 | -105.4146
15.2686 | -105.4142
15.2724 | -105.4152
15.2761 | -105.4142
15.2799 | -105.4119
15.2836 | -105.4108
15.2874 | -105.4101
15.2911 | -105.406
15.2949 | -105.3999
15.2986 | -105.3958
15.3024 | -105.3918
15.3061 | -105.3876
15.3099 | -105.3822
15.3136 | -105.3736
15.3174 | -105.3701
15.3211 | -105.3677
15.3249 | -105.3655
15.3286 | -105.361
15.3324 | -105.3549
15.3361 | -105.3469
15.3399 | -105.339
15.3436 | -105.3304
15.3474 | -105.3213
15.3511 | -105.3146
15.3549 | -105.3085
15.3587 | -105.3014
15.3624 | -105.2925
15.3662 | -105.2857
15.3699 | -105.2845
15.3737 | -105.2815
15.3774 | -105.2796
15.3812 | -105.2785
15.3849 | -105.2787
15.3887 | -105.2801
15.3924 | -105.2829
15.3962 | -105.2887
15.3999 | -105.2967
15.4037 | -105.3072
15.4074 | -105.3183
15.4112 | -105.3309
15.4149 | -105.3448
15.4187 | -105.3579
15.4224 | -105.3715
15.4262 | -105.3851
15.4299 | -105.4046
15.4337 | -105.4228
15.4374 | -105.4404
15.4412 | -105.4568
15.4449 | -105.4707
15.4487 | -105.4827
15.4524 | -105.4904
15.4562 | -105.4981
15.4599 | -105.5039
15.4637 | -105.5094
15.4674 | -105.5118
15.4712 | -105.5121
15.4749 | -105.51
15.4787 | -105.5064
15.4824 | -105.5005
15.4862 | -105.4898
15.4899 | -105.4781
15.4937 | -105.4627
15.4974 | -105.4458
15.5012 | -105.4251
15.5049 | -105.4024
15.5087 | -105.3788
15.5124 | -105.3547
15.5162 | -105.3303
15.5199 | -105.3045
15.5237 | -105.2789
15.5274 | -105.2544
15.5312 | -105.2314
15.5349 | -105.2109
15.5387 | -105.1903
15.5424 | -105.1704
15.5462 | -105.153
15.5499 | -105.1385
15.5537 | -105.1268
15.5574 | -105.1188
15.5612 | -105.1138
15.5649 | -105.1119
15.5687 | -105.1134
15.5724 | -105.1159
15.5762 | -105.1226
15.5799 | -105.1318
15.5837 | -105.1411
15.5874 | -105.155
15.5912 | -105.1671
15.5949 | -105.1816
15.5987 | -105.1975
15.6024 | -105.2166
15.6062 | -105.2366
15.6099 | -105.2571
15.6137 | -105.276
15.6174 | -105.2934
15.6212 | -105.3097
15.6249 | -105.3219
15.6287 | -105.3308
15.6324 | -105.3401
15.6362 | -105.349
15.6399 | -105.3575
15.6437 | -105.3615
15.6474 | -105.3625
15.6512 | -105.361
15.6549 | -105.3591
15.6587 | -105.3527
15.6624 | -105.3413
15.6662 | -105.3307
15.6699 | -105.3178
15.6737 | -105.3037
15.6775 | -105.2891
15.6812 | -105.2743
15.685 | -105.2563
15.6887 | -105.2341
15.6925 | -105.2152
15.6962 | -105.1957
15.7 | -105.1775
15.7037 | -105.1572
15.7075 | -105.137
15.7112 | -105.1187
15.715 | -105.1016
15.7187 | -105.0877
15.7225 | -105.075
15.7262 | -105.0647
15.73 | -105.0531
15.7337 | -105.0441
15.7375 | -105.0371
15.7412 | -105.0304
15.745 | -105.0227
15.7487 | -105.0148
15.7525 | -105.01
15.7562 | -105.0062
15.76 | -105.0014
15.7637 | -104.9984
15.7675 | -104.9949
15.7712 | -104.9904
15.775 | -104.9889
15.7787 | -104.9898
15.7825 | -104.9902
15.7862 | -104.9916
15.79 | -104.9939
15.7937 | -104.9946
15.7975 | -104.9942
15.8012 | -104.9907
15.805 | -104.9878
15.8087 | -104.9911
15.8125 | -104.9914
15.8162 | -104.9894
15.82 | -104.9869
15.8237 | -104.9833
15.8275 | -104.9772
15.8312 | -104.9702
15.835 | -104.9625
15.8387 | -104.9535
15.8425 | -104.9449
15.8462 | -104.9379
15.85 | -104.9306
15.8537 | -104.9218
15.8575 | -104.9109
15.8612 | -104.8987
15.865 | -104.8863
15.8687 | -104.876
15.8725 | -104.869
15.8762 | -104.8623
15.88 | -104.8572
15.8837 | -104.8571
15.8875 | -104.8595
15.8912 | -104.8593
15.895 | -104.8644
15.8987 | -104.8718
15.9025 | -104.8778
15.9062 | -104.8857
15.91 | -104.8978
15.9137 | -104.91
15.9175 | -104.9229
15.9212 | -104.9376
15.925 | -104.9573
15.9287 | -104.9775
15.9325 | -104.9959
15.9362 | -105.015
15.94 | -105.0348
15.9437 | -105.052
15.9475 | -105.0695
15.9512 | -105.0878
15.955 | -105.1028
15.9587 | -105.1131
15.9625 | -105.1228
15.9662 | -105.1326
15.97 | -105.1398
15.9737 | -105.1431
15.9775 | -105.1425
15.9812 | -105.1396
15.985 | -105.1347
15.9887 | -105.1278
15.9925 | -105.1165
15.9962 | -105.1026
16 | -105.0826
16.0038 | -105.0597
16.0075 | -105.0343
16.0113 | -105.0049
16.015 | -104.9745
16.0188 | -104.9385
16.0225 | -104.9064
16.0263 | -104.8772
16.03 | -104.8468
16.0338 | -104.8125
16.0375 | -104.7766
16.0413 | -104.7416
16.045 | -104.7099
16.0488 | -104.6792
16.0525 | -104.6513
16.0563 | -104.6255
16.06 | -104.6012
16.0638 | -104.5786
16.0675 | -104.5588
16.0713 | -104.5424
16.075 | -104.5262
16.0788 | -104.5116
16.0825 | -104.5031
16.0863 | -104.4979
16.09 | -104.4927
16.0938 | -104.4882
16.0975 | -104.4853
16.1013 | -104.4808
16.105 | -104.4752
16.1088 | -104.4738
16.1125 | -104.4744
16.1163 | -104.477
16.12 | -104.4806
16.1238 | -104.4831
16.1275 | -104.4843
16.1313 | -104.4869
16.135 | -104.4904
16.1388 | -104.4896
16.1425 | -104.4885
16.1463 | -104.4852
16.15 | -104.4785
16.1538 | -104.4673
16.1575 | -104.4537
16.1613 | -104.4368
16.165 | -104.4185
16.1688 | -104.4002
16.1725 | -104.3826
16.1763 | -104.3589
16.18 | -104.3314
16.1838 | -104.3061
16.1875 | -104.2845
16.1913 | -104.2563
16.195 | -104.223
16.1988 | -104.1928
16.2025 | -104.1634
16.2063 | -104.1361
16.21 | -104.1063
16.2138 | -104.076
16.2175 | -104.0507
16.2213 | -104.0264
16.225 | -104
16.2288 | -103.9781
16.2325 | -103.9645
16.2363 | -103.9461
16.24 | -103.9293
16.2438 | -103.9148
16.2475 | -103.9002
16.2513 | -103.8888
16.255 | -103.8735
16.2588 | -103.8585
16.2625 | -103.8478
16.2663 | -103.839
16.27 | -103.8317
16.2738 | -103.8262
16.2775 | -103.8216
16.2813 | -103.8175
16.285 | -103.8145
16.2888 | -103.813
16.2925 | -103.813
16.2963 | -103.8152
16.3 | -103.8195
16.3038 | -103.8238
16.3075 | -103.831
16.3113 | -103.8422
16.315 | -103.853
16.3188 | -103.8625
16.3225 | -103.8739
16.3263 | -103.8877
16.3301 | -103.902
16.3338 | -103.9204
16.3376 | -103.9364
16.3413 | -103.948
16.3451 | -103.9585
16.3488 | -103.9688
16.3526 | -103.9775
16.3563 | -103.9868
16.3601 | -103.9937
16.3638 | -103.9993
16.3676 | -104.0034
16.3713 | -104.0068
16.3751 | -104.0095
16.3788 | -104.0095
16.3826 | -104.0077
16.3863 | -104.005
16.3901 | -104
16.3938 | -103.9958
16.3976 | -103.9874
16.4013 | -103.9773
16.4051 | -103.9664
16.4088 | -103.9574
16.4126 | -103.9491
16.4163 | -103.938
16.4201 | -103.927
16.4238 | -103.9178
16.4276 | -103.9104
16.4313 | -103.9013
16.4351 | -103.8932
16.4388 | -103.8852
16.4426 | -103.88
16.4463 | -103.8759
16.4501 | -103.8729
16.4538 | -103.8743
16.4576 | -103.8815
16.4613 | -103.8902
16.4651 | -103.8992
16.4688 | -103.91
16.4726 | -103.9247
16.4763 | -103.9429
16.4801 | -103.96
16.4838 | -103.9797
16.4876 | -104.0014
16.4913 | -104.0231
16.4951 | -104.0459
16.4988 | -104.068
16.5026 | -104.091
16.5063 | -104.1129
16.5101 | -104.1336
16.5138 | -104.1537
16.5176 | -104.171
16.5213 | -104.1865
16.5251 | -104.2005
16.5288 | -104.214
16.5326 | -104.2269
16.5363 | -104.2369
16.5401 | -104.246
16.5438 | -104.2533
16.5476 | -104.2584
16.5513 | -104.2605
16.5551 | -104.261
16.5588 | -104.2644
16.5626 | -104.2669
16.5663 | -104.2684
16.5701 | -104.2658
16.5738 | -104.2671
16.5776 | -104.2687
16.5813 | -104.268
16.5851 | -104.2647
16.5888 | -104.2653
16.5926 | -104.2635
16.5963 | -104.2612
16.6001 | -104.2578
16.6038 | -104.2536
16.6076 | -104.2496
16.6113 | -104.2483
16.6151 | -104.2465
16.6188 | -104.2436
16.6226 | -104.2411
16.6263 | -104.2397
16.6301 | -104.2382
16.6338 | -104.237
16.6376 | -104.2342
16.6413 | -104.2323
16.6451 | -104.2339
16.6489 | -104.2364
16.6526 | -104.2391
16.6564 | -104.2429
16.6601 | -104.2505
16.6639 | -104.2581
16.6676 | -104.2664
16.6714 | -104.2773
16.6751 | -104.2886
16.6789 | -104.3014
16.6826 | -104.3169
16.6864 | -104.3323
16.6901 | -104.348
16.6939 | -104.3653
16.6976 | -104.3822
16.7014 | -104.4008
16.7051 | -104.4239
16.7089 | -104.4475
16.7126 | -104.4693
16.7164 | -104.494
16.7201 | -104.5189
16.7239 | -104.5444
16.7276 | -104.5712
16.7314 | -104.5974
16.7351 | -104.6224
16.7389 | -104.6444
16.7426 | -104.6631
16.7464 | -104.6785
16.7501 | -104.6912
16.7539 | -104.7028
16.7576 | -104.7107
16.7614 | -104.7115
16.7651 | -104.7078
16.7689 | -104.7013
16.7726 | -104.6694
16.7764 | -104.6542
16.7801 | -104.6347
16.7839 | -104.6119
16.7876 | -104.5865
16.7914 | -104.5591
16.7951 | -104.5313
16.7989 | -104.4999
16.8026 | -104.4663
16.8064 | -104.4279
16.8101 | -104.3902
16.8139 | -104.353
16.8176 | -104.3162
16.8214 | -104.2812
16.8251 | -104.2471
16.8289 | -104.2104
16.8326 | -104.1743
16.8364 | -104.141
16.8401 | -104.1099
16.8439 | -104.0803
16.8476 | -104.0541
16.8514 | -104.0283
16.8551 | -104.0037
16.8589 | -103.9819
16.8626 | -103.9621
16.8664 | -103.9479
16.8701 | -103.9371
16.8739 | -103.9306
16.8776 | -103.926
16.8814 | -103.9225
16.8851 | -103.9195
16.8889 | -103.9196
16.8926 | -103.9227
16.8964 | -103.9281
16.9001 | -103.9353
16.9039 | -103.9462
16.9076 | -103.9572
16.9114 | -103.9672
16.9151 | -103.9792
16.9189 | -103.992
16.9226 | -104.0019
16.9264 | -104.0121
16.9301 | -104.0229
16.9339 | -104.0336
16.9376 | -104.0427
16.9414 | -104.05
16.9451 | -104.0564
16.9489 | -104.0621
16.9526 | -104.0677
16.9564 | -104.0711
16.9601 | -104.0713
16.9639 | -104.0661
16.9677 | -104.0584
16.9714 | -104.0496
16.9752 | -104.0332
16.9789 | -104.0151
16.9827 | -103.9948
16.9864 | -103.9725
16.9902 | -103.9481
16.9939 | -103.9162
16.9977 | -103.877
17.0014 | -103.8377
17.0052 | -103.7956
17.0089 | -103.7474
17.0127 | -103.6969
17.0164 | -103.6453
17.0202 | -103.5923
17.0239 | -103.5371
17.0277 | -103.4778
17.0314 | -103.4175
17.0352 | -103.3573
17.0389 | -103.2959
17.0427 | -103.2326
17.0464 | -103.1675
17.0502 | -103.1034
17.0539 | -103.039
17.0577 | -102.9764
17.0614 | -102.9153
17.0652 | -102.857
17.0689 | -102.8013
17.0727 | -102.7498
17.0764 | -102.7028
17.0802 | -102.6605
17.0839 | -102.6218
17.0877 | -102.5861
17.0914 | -102.5537
17.0952 | -102.5276
17.0989 | -102.5086
17.1027 | -102.4955
17.1064 | -102.4884
17.1102 | -102.4831
17.1139 | -102.4828
17.1177 | -102.4887
17.1214 | -102.5016
17.1252 | -102.5208
17.1289 | -102.5435
17.1327 | -102.5692
17.1364 | -102.5999
17.1402 | -102.6341
17.1439 | -102.6718
17.1477 | -102.7135
17.1514 | -102.7589
17.1552 | -102.8092
17.1589 | -102.8598
17.1627 | -102.9114
17.1664 | -102.966
17.1702 | -103.0211
17.1739 | -103.0758
17.1777 | -103.1324
17.1814 | -103.19
17.1852 | -103.2459
17.1889 | -103.3006
17.1927 | -103.3536
17.1964 | -103.4053
17.2002 | -103.4579
17.2039 | -103.5085
17.2077 | -103.5593
17.2114 | -103.6069
17.2152 | -103.6486
17.2189 | -103.6892
17.2227 | -103.7276
17.2264 | -103.7613
17.2302 | -103.7895
17.2339 | -103.8108
17.2377 | -103.8317
17.2414 | -103.8513
17.2452 | -103.8621
17.2489 | -103.8633
17.2527 | -103.8592
17.2564 | -103.8484
17.2602 | -103.8329
17.2639 | -103.8144
17.2677 | -103.7916
17.2714 | -103.7633
17.2752 | -103.7295
17.2789 | -103.6914
17.2827 | -103.649
17.2865 | -103.6045
17.2902 | -103.5572
17.294 | -103.5089
17.2977 | -103.4621
17.3015 | -103.4089
17.3052 | -103.356
17.309 | -103.3028
17.3127 | -103.2393
17.3165 | -103.1841
17.3202 | -103.1297
17.324 | -103.0715
17.3277 | -103.0132
17.3315 | -102.9554
17.3352 | -102.9013
17.339 | -102.8475
17.3427 | -102.793
17.3465 | -102.7345
17.3502 | -102.6767
17.354 | -102.6195
17.3577 | -102.5651
17.3615 | -102.5093
17.3652 | -102.4513
17.369 | -102.3966
17.3727 | -102.344
17.3765 | -102.2964
17.3802 | -102.2515
17.384 | -102.2057
17.3877 | -102.1603
17.3915 | -102.1186
17.3952 | -102.0787
17.399 | -102.0428
17.4027 | -102.0091
17.4065 | -101.9783
17.4102 | -101.9513
17.414 | -101.9279
17.4177 | -101.9083
17.4215 | -101.8934
17.4252 | -101.88
17.429 | -101.8708
17.4327 | -101.8646
17.4365 | -101.8624
17.4402 | -101.8641
17.444 | -101.8692
17.4477 | -101.8787
17.4515 | -101.8926
17.4552 | -101.9092
17.459 | -101.9301
17.4627 | -101.9549
17.4665 | -101.9844
17.4702 | -102.0168
17.474 | -102.0511
17.4777 | -102.0893
17.4815 | -102.1318
17.4852 | -102.1777
17.489 | -102.2234
17.4927 | -102.2701
17.4965 | -102.3179
17.5002 | -102.3662
17.504 | -102.4141
17.5077 | -102.4597
17.5115 | -102.5034
17.5152 | -102.544
17.519 | -102.5831
17.5227 | -102.617
17.5265 | -102.6467
17.5302 | -102.6726
17.534 | -102.6918
17.5377 | -102.703
17.5415 | -102.7078
17.5452 | -102.7003
17.549 | -102.6857
17.5527 | -102.663
17.5565 | -102.6328
17.5602 | -102.5962
17.564 | -102.5487
17.5677 | -102.4932
17.5715 | -102.4301
17.5752 | -102.3632
17.579 | -102.29
17.5827 | -102.2104
17.5865 | -102.1271
17.5902 | -102.0427
17.594 | -101.9534
17.5978 | -101.8652
17.6015 | -101.7748
17.6053 | -101.684
17.609 | -101.5914
17.6128 | -101.5002
17.6165 | -101.411
17.6203 | -101.324
17.624 | -101.2382
17.6278 | -101.1535
17.6315 | -101.0726
17.6353 | -100.9967
17.639 | -100.9246
17.6428 | -100.857
17.6465 | -100.794
17.6503 | -100.7348
17.654 | -100.6802
17.6578 | -100.6321
17.6615 | -100.5893
17.6653 | -100.5525
17.669 | -100.5212
17.6728 | -100.496
17.6765 | -100.479
17.6803 | -100.4667
17.684 | -100.4617
17.6878 | -100.4627
17.6915 | -100.4703
17.6953 | -100.4852
17.699 | -100.5046
17.7028 | -100.5312
17.7065 | -100.5642
17.7103 | -100.6028
17.714 | -100.6459
17.7178 | -100.6934
17.7215 | -100.7446
17.7253 | -100.8013
17.729 | -100.861
17.7328 | -100.9227
17.7365 | -100.9874
17.7403 | -101.0558
17.744 | -101.1271
17.7478 | -101.1992
17.7515 | -101.2715
17.7553 | -101.3433
17.759 | -101.4152
17.7628 | -101.4873
17.7665 | -101.557
17.7703 | -101.6233
17.774 | -101.6875
17.7778 | -101.7473
17.7815 | -101.8011
17.7853 | -101.8503
17.789 | -101.895
17.7928 | -101.9307
17.7965 | -101.9625
17.8003 | -101.9886
17.804 | -102.0055
17.8078 | -102.0136
17.8115 | -102.0149
17.8153 | -102.01
17.819 | -101.9954
17.8228 | -101.9729
17.8265 | -101.9446
17.8303 | -101.9071
17.834 | -101.8639
17.8378 | -101.8119
17.8415 | -101.7543
17.8453 | -101.6941
17.849 | -101.627
17.8528 | -101.5404
17.8565 | -101.4742
17.8603 | -101.406
17.864 | -101.3366
17.8678 | -101.2663
17.8715 | -101.1968
17.8753 | -101.1258
17.879 | -101.0564
17.8828 | -100.9882
17.8865 | -100.9219
17.8903 | -100.8588
17.894 | -100.7986
17.8978 | -100.7403
17.9015 | -100.6858
17.9053 | -100.6356
17.909 | -100.5857
17.9128 | -100.5389
17.9165 | -100.4982
17.9203 | -100.4605
17.9241 | -100.4263
17.9278 | -100.3944
17.9316 | -100.3647
17.9353 | -100.3409
17.9391 | -100.3185
17.9428 | -100.3008
17.9466 | -100.2872
17.9503 | -100.2774
17.9541 | -100.2713
17.9578 | -100.2679
17.9616 | -100.2686
17.9653 | -100.2731
17.9691 | -100.283
17.9728 | -100.2967
17.9766 | -100.3145
17.9803 | -100.3357
17.9841 | -100.3604
17.9878 | -100.3886
17.9916 | -100.4185
17.9953 | -100.4511
17.9991 | -100.4864
18.0028 | -100.5255
18.0066 | -100.5675
18.0103 | -100.6121
18.0141 | -100.6576
18.0178 | -100.7037
18.0216 | -100.7508
18.0253 | -100.7998
18.0291 | -100.8504
18.0328 | -100.9032
18.0366 | -100.9556
18.0403 | -101.0091
18.0441 | -101.0634
18.0478 | -101.1195
18.0516 | -101.1746
18.0553 | -101.2266
18.0591 | -101.2788
18.0628 | -101.3291
18.0666 | -101.3776
18.0703 | -101.4259
18.0741 | -101.4721
18.0778 | -101.513
18.0816 | -101.5494
18.0853 | -101.5848
18.0891 | -101.6161
18.0928 | -101.6447
18.0966 | -101.6678
18.1003 | -101.6857
18.1041 | -101.7023
18.1078 | -101.711
18.1116 | -101.716
18.1153 | -101.7156
18.1191 | -101.71
18.1228 | -101.7012
18.1266 | -101.6892
18.1303 | -101.6735
18.1341 | -101.6546
18.1378 | -101.6316
18.1416 | -101.607
18.1453 | -101.5789
18.1491 | -101.5477
18.1528 | -101.5155
18.1566 | -101.4811
18.1603 | -101.4473
18.1641 | -101.4113
18.1678 | -101.3795
18.1716 | -101.3482
18.1753 | -101.3133
18.1791 | -101.28
18.1828 | -101.2482
18.1866 | -101.2193
18.1903 | -101.1923
18.1941 | -101.1644
18.1978 | -101.1401
18.2016 | -101.114
18.2053 | -101.0907
18.2091 | -101.0686
18.2128 | -101.0483
18.2166 | -101.0281
18.2203 | -101.0109
18.2241 | -100.9953
18.2278 | -100.9805
18.2316 | -100.9692
18.2353 | -100.9602
18.2391 | -100.9526
18.2429 | -100.9476
18.2466 | -100.9463
18.2504 | -100.9474
18.2541 | -100.9504
18.2579 | -100.9543
18.2616 | -100.9629
18.2654 | -100.9747
18.2691 | -100.9886
18.2729 | -101.0059
18.2766 | -101.0252
18.2804 | -101.0458
18.2841 | -101.0686
18.2879 | -101.0919
18.2916 | -101.1146
18.2954 | -101.1429
18.2991 | -101.1716
18.3029 | -101.201
18.3066 | -101.2301
18.3104 | -101.2596
18.3141 | -101.2903
18.3179 | -101.3187
18.3216 | -101.3474
18.3254 | -101.3757
18.3291 | -101.404
18.3329 | -101.4319
18.3366 | -101.4604
18.3404 | -101.4882
18.3441 | -101.5151
18.3479 | -101.5404
18.3516 | -101.5638
18.3554 | -101.5867
18.3591 | -101.6104
18.3629 | -101.6327
18.3666 | -101.6542
18.3704 | -101.6745
18.3741 | -101.6953
18.3779 | -101.715
18.3816 | -101.7342
18.3854 | -101.7509
18.3891 | -101.7673
18.3929 | -101.7831
18.3966 | -101.7967
18.4004 | -101.8117
18.4041 | -101.8242
18.4079 | -101.8383
18.4116 | -101.8531
18.4154 | -101.8677
18.4191 | -101.882
18.4229 | -101.8926
18.4266 | -101.9044
18.4304 | -101.9163
18.4341 | -101.9253
18.4379 | -101.9343
18.4416 | -101.9387
18.4454 | -101.9395
18.4491 | -101.9396
18.4529 | -101.9402
18.4566 | -101.9397
18.4604 | -101.9368
18.4641 | -101.932
18.4679 | -101.9246
18.4716 | -101.9197
18.4754 | -101.9132
18.4791 | -101.905
18.4829 | -101.8922
18.4866 | -101.8783
18.4904 | -101.8654
18.4941 | -101.8506
18.4979 | -101.8337
18.5016 | -101.8152
18.5054 | -101.7965
18.5091 | -101.7762
18.5129 | -101.7546
18.5166 | -101.7342
18.5204 | -101.7156
18.5241 | -101.6928
18.5279 | -101.6669
18.5316 | -101.6417
18.5354 | -101.6169
18.5391 | -101.5937
18.5429 | -101.5718
18.5466 | -101.5502
18.5504 | -101.5305
18.5541 | -101.5131
18.5579 | -101.4942
18.5617 | -101.4766
18.5654 | -101.4611
18.5692 | -101.448
18.5729 | -101.4355
18.5767 | -101.4289
18.5804 | -101.4259
18.5842 | -101.4239
18.5879 | -101.4216
18.5917 | -101.4212
18.5954 | -101.4235
18.5992 | -101.4282
18.6029 | -101.434
18.6067 | -101.4431
18.6104 | -101.4527
18.6142 | -101.4654
18.6179 | -101.4803
18.6217 | -101.4951
18.6254 | -101.5114
18.6292 | -101.5265
18.6329 | -101.5415
18.6367 | -101.5595
18.6404 | -101.5765
18.6442 | -101.5937
18.6479 | -101.6099
18.6517 | -101.6235
18.6554 | -101.636
18.6592 | -101.6492
18.6629 | -101.659
18.6667 | -101.6715
18.6704 | -101.6836
18.6742 | -101.6896
18.6779 | -101.6917
18.6817 | -101.6931
18.6854 | -101.6909
18.6892 | -101.687
18.6929 | -101.6814
18.6967 | -101.6725
18.7004 | -101.6636
18.7042 | -101.6524
18.7079 | -101.6423
18.7117 | -101.6325
18.7154 | -101.618
18.7192 | -101.603
18.7229 | -101.5867
18.7267 | -101.5698
18.7304 | -101.5516
18.7342 | -101.5315
18.7379 | -101.5091
18.7417 | -101.4886
18.7454 | -101.4685
18.7492 | -101.45
18.7529 | -101.4327
18.7567 | -101.4157
18.7604 | -101.3991
18.7642 | -101.3831
18.7679 | -101.3701
18.7717 | -101.3592
18.7754 | -101.35
18.7792 | -101.3445
18.7829 | -101.3382
18.7867 | -101.3343
18.7904 | -101.3348
18.7942 | -101.3417
18.7979 | -101.3517
18.8017 | -101.3628
18.8054 | -101.3781
18.8092 | -101.3967
18.8129 | -101.417
18.8167 | -101.4388
18.8204 | -101.4623
18.8242 | -101.489
18.8279 | -101.5184
18.8317 | -101.5522
18.8354 | -101.5856
18.8392 | -101.6214
18.8429 | -101.6576
18.8467 | -101.6945
18.8504 | -101.7328
18.8542 | -101.7715
18.8579 | -101.8078
18.8617 | -101.8428
18.8654 | -101.8735
18.8692 | -101.9057
18.8729 | -101.9374
18.8767 | -101.9688
18.8804 | -101.9965
18.8842 | -102.022
18.888 | -102.0446
18.8917 | -102.0646
18.8955 | -102.0836
18.8992 | -102.1002
18.903 | -102.109
18.9067 | -102.1145
18.9105 | -102.116
18.9142 | -102.1157
18.918 | -102.1132
18.9217 | -102.1038
18.9255 | -102.0893
18.9292 | -102.0728
18.933 | -102.0532
18.9367 | -102.0205
18.9405 | -101.9999
18.9442 | -101.9819
18.948 | -101.9609
18.9517 | -101.9344
18.9555 | -101.911
18.9592 | -101.888
18.963 | -101.862
18.9667 | -101.8353
18.9705 | -101.8085
18.9742 | -101.7814
18.978 | -101.754
18.9817 | -101.7297
18.9855 | -101.708
18.9892 | -101.6834
18.993 | -101.6638
18.9967 | -101.646
19.0005 | -101.6307
19.0042 | -101.6187
19.008 | -101.6087
19.0117 | -101.5993
19.0155 | -101.5943
19.0192 | -101.5918
19.023 | -101.5912
19.0267 | -101.5921
19.0305 | -101.5958
19.0342 | -101.6067
19.038 | -101.6197
19.0417 | -101.6315
19.0455 | -101.6479
19.0492 | -101.667
19.053 | -101.6909
19.0567 | -101.715
19.0605 | -101.7377
19.0642 | -101.7597
19.068 | -101.7811
19.0717 | -101.8031
19.0755 | -101.8231
19.0792 | -101.8449
19.083 | -101.8665
19.0867 | -101.8845
19.0905 | -101.9043
19.0942 | -101.9226
19.098 | -101.9351
19.1017 | -101.9427
19.1055 | -101.9521
19.1092 | -101.9598
19.113 | -101.96
19.1167 | -101.9564
19.1205 | -101.9515
19.1242 | -101.9439
19.128 | -101.9314
19.1317 | -101.9141
19.1355 | -101.8922
19.1392 | -101.8626
19.143 | -101.8317
19.1467 | -101.7989
19.1505 | -101.7626
19.1542 | -101.7247
19.158 | -101.6862
19.1617 | -101.6452
19.1655 | -101.6034
19.1692 | -101.5603
19.173 | -101.5195
19.1767 | -101.4793
19.1805 | -101.4372
19.1842 | -101.392
19.188 | -101.3491
19.1917 | -101.3103
19.1955 | -101.272
19.1992 | -101.2342
19.203 | -101.1969
19.2068 | -101.1607
19.2105 | -101.1292
19.2143 | -101.1002
19.218 | -101.0721
19.2218 | -101.046
19.2255 | -101.0247
19.2293 | -101.0067
19.233 | -100.9903
19.2368 | -100.9783
19.2405 | -100.968
19.2443 | -100.9595
19.248 | -100.9551
19.2518 | -100.954
19.2555 | -100.9513
19.2593 | -100.9521
19.263 | -100.9567
19.2668 | -100.9598
19.2705 | -100.9658
19.2743 | -100.975
19.278 | -100.9835
19.2818 | -100.9942
19.2855 | -101.0064
19.2893 | -101.0182
19.293 | -101.0283
19.2968 | -101.0381
19.3005 | -101.0459
19.3043 | -101.0551
19.308 | -101.0641
19.3118 | -101.0725
19.3155 | -101.0836
19.3193 | -101.0901
19.323 | -101.0945
19.3268 | -101.1009
19.3305 | -101.107
19.3343 | -101.1127
19.338 | -101.1166
19.3418 | -101.1213
19.3455 | -101.1244
19.3493 | -101.1271
19.353 | -101.1265
19.3568 | -101.1264
19.3605 | -101.1271
19.3643 | -101.1238
19.368 | -101.1207
19.3718 | -101.1223
19.3755 | -101.1228
19.3793 | -101.1198
19.383 | -101.1168
19.3868 | -101.1158
19.3905 | -101.1165
19.3943 | -101.1158
19.398 | -101.1189
19.4018 | -101.1237
19.4055 | -101.1292
19.4093 | -101.1377
19.413 | -101.1487
19.4168 | -101.1545
19.4205 | -101.1624
19.4243 | -101.1747
19.428 | -101.1866
19.4318 | -101.1999
19.4355 | -101.2146
19.4393 | -101.2314
19.443 | -101.2499
19.4468 | -101.2689
19.4505 | -101.2886
19.4543 | -101.3105
19.458 | -101.3346
19.4618 | -101.3536
19.4655 | -101.3743
19.4693 | -101.3966
19.473 | -101.4197
19.4768 | -101.4413
19.4805 | -101.4616
19.4843 | -101.4818
19.488 | -101.5022
19.4918 | -101.525
19.4955 | -101.5478
19.4993 | -101.572
19.503 | -101.5965
19.5068 | -101.6164
19.5105 | -101.6374
19.5143 | -101.6578
19.518 | -101.6781
19.5218 | -101.6991
19.5256 | -101.718
19.5293 | -101.7358
19.5331 | -101.7556
19.5368 | -101.7739
19.5406 | -101.7912
19.5443 | -101.8118
19.5481 | -101.8304
19.5518 | -101.8464
19.5556 | -101.8636
19.5593 | -101.8797
19.5631 | -101.8962
19.5668 | -101.9129
19.5706 | -101.9311
19.5743 | -101.9494
19.5781 | -101.9684
19.5818 | -101.9846
19.5856 | -101.9978
19.5893 | -102.0119
19.5931 | -102.0235
19.5968 | -102.0371
19.6006 | -102.0513
19.6043 | -102.0618
19.6081 | -102.0736
19.6118 | -102.0843
19.6156 | -102.0957
19.6193 | -102.1075
19.6231 | -102.1195
19.6268 | -102.131
19.6306 | -102.1448
19.6343 | -102.1546
19.6381 | -102.1645
19.6418 | -102.1753
19.6456 | -102.186
19.6493 | -102.1967
19.6531 | -102.2089
19.6568 | -102.2238
19.6606 | -102.2395
19.6643 | -102.2559
19.6681 | -102.2716
19.6718 | -102.2882
19.6756 | -102.3015
19.6793 | -102.317
19.6831 | -102.3347
19.6868 | -102.3511
19.6906 | -102.367
19.6943 | -102.3854
19.6981 | -102.4017
19.7018 | -102.4175
19.7056 | -102.4345
19.7093 | -102.4505
19.7131 | -102.467
19.7168 | -102.4849
19.7206 | -102.5021
19.7243 | -102.5178
19.7281 | -102.5344
19.7318 | -102.5531
19.7356 | -102.5739
19.7393 | -102.5973
19.7431 | -102.6153
19.7468 | -102.6346
19.7506 | -102.6545
19.7543 | -102.6721
19.7581 | -102.692
19.7618 | -102.714
19.7656 | -102.7328
19.7693 | -102.7497
19.7731 | -102.7675
19.7768 | -102.7864
19.7806 | -102.8048
19.7843 | -102.8214
19.7881 | -102.841
19.7918 | -102.8597
19.7956 | -102.878
19.7993 | -102.899
19.8031 | -102.9154
19.8068 | -102.9328
19.8106 | -102.9505
19.8143 | -102.9666
19.8181 | -102.98
19.8218 | -102.9916
19.8256 | -102.9996
19.8293 | -103.0068
19.8331 | -103.0138
19.8368 | -103.0215
19.8406 | -103.0275
19.8444 | -103.0338
19.8481 | -103.039
19.8519 | -103.043
19.8556 | -103.0468
19.8594 | -103.0492
19.8631 | -103.0508
19.8669 | -103.0513
19.8706 | -103.0497
19.8744 | -103.0489
19.8781 | -103.0488
19.8819 | -103.0492
19.8856 | -103.0493
19.8894 | -103.049
19.8931 | -103.0502
19.8969 | -103.05
19.9006 | -103.0502
19.9044 | -103.0517
19.9081 | -103.0521
19.9119 | -103.0527
19.9156 | -103.0552
19.9194 | -103.0567
19.9231 | -103.0573
19.9269 | -103.0584
19.9306 | -103.0577
19.9344 | -103.0581
19.9381 | -103.0579
19.9419 | -103.0581
19.9456 | -103.0579
19.9494 | -103.0574
19.9531 | -103.058
19.9569 | -103.0585
19.9606 | -103.0578
19.9644 | -103.0579
19.9681 | -103.0584
19.9719 | -103.0574
19.9756 | -103.0571
19.9794 | -103.059
19.9831 | -103.06
19.9869 | -103.0593
19.9906 | -103.0607
19.9944 | -103.0608
19.9981 | -103.0582
20.0019 | -103.058
20.0056 | -103.0575
20.0094 | -103.0564
20.0131 | -103.0548
20.0169 | -103.0548
20.0206 | -103.0543
20.0244 | -103.0534
20.0281 | -103.0553
20.0319 | -103.0543
20.0356 | -103.054
20.0394 | -103.0547
20.0431 | -103.0543
20.0469 | -103.0511
20.0506 | -103.0504
20.0544 | -103.0471
20.0581 | -103.0403
20.0619 | -103.034
20.0656 | -103.0304
20.0694 | -103.0253
20.0731 | -103.019
20.0769 | -103.0127
20.0806 | -103.0074
20.0844 | -102.9987
20.0881 | -102.9926
20.0919 | -102.985
20.0956 | -102.975
20.0994 | -102.9663
20.1031 | -102.9555
20.1069 | -102.9452
20.1106 | -102.9377
20.1144 | -102.9302
20.1181 | -102.9206
20.1219 | -102.9091
20.1256 | -102.8999
20.1294 | -102.8914
20.1331 | -102.8836
20.1369 | -102.8761
20.1406 | -102.8684
20.1444 | -102.8625
20.1481 | -102.8582
20.1519 | -102.8536
20.1556 | -102.8503
20.1594 | -102.848
20.1632 | -102.8452
20.1669 | -102.8438
20.1707 | -102.8416
20.1744 | -102.8407
20.1782 | -102.8427
20.1819 | -102.8459
20.1857 | -102.8491
20.1894 | -102.8523
20.1932 | -102.8535
20.1969 | -102.8572
20.2007 | -102.8649
20.2044 | -102.8725
20.2082 | -102.8797
20.2119 | -102.8848
20.2157 | -102.8926
20.2194 | -102.8994
20.2232 | -102.9075
20.2269 | -102.9158
20.2307 | -102.9227
20.2344 | -102.9303
20.2382 | -102.9393
20.2419 | -102.947
20.2457 | -102.9534
20.2494 | -102.9599
20.2532 | -102.966
20.2569 | -102.971
20.2607 | -102.9768
20.2644 | -102.9825
20.2682 | -102.9848
20.2719 | -102.9889
20.2757 | -102.9917
20.2794 | -102.9956
20.2832 | -103
20.2869 | -103.0033
20.2907 | -103.0087
20.2944 | -103.014
20.2982 | -103.0175
20.3019 | -103.0201
20.3057 | -103.0235
20.3094 | -103.0269
20.3132 | -103.0315
20.3169 | -103.0344
20.3207 | -103.0375
20.3244 | -103.0426
20.3282 | -103.0483
20.3319 | -103.0532
20.3357 | -103.059
20.3394 | -103.066
20.3432 | -103.0743
20.3469 | -103.082
20.3507 | -103.0909
20.3544 | -103.0973
20.3582 | -103.1038
20.3619 | -103.1121
20.3657 | -103.122
20.3694 | -103.1292
20.3732 | -103.137
20.3769 | -103.1442
20.3807 | -103.1517
20.3844 | -103.1599
20.3882 | -103.1662
20.3919 | -103.1733
20.3957 | -103.1788
20.3994 | -103.1819
20.4032 | -103.1858
20.4069 | -103.1902
20.4107 | -103.1945
20.4144 | -103.1949
20.4182 | -103.1928
20.4219 | -103.1913
20.4257 | -103.1902
20.4294 | -103.1892
20.4332 | -103.1874
20.4369 | -103.1852
20.4407 | -103.1833
20.4444 | -103.1814
20.4482 | -103.1746
20.4519 | -103.1675
20.4557 | -103.1582
20.4594 | -103.1498
20.4632 | -103.1411
20.4669 | -103.1314
20.4707 | -103.1213
20.4744 | -103.1098
20.4782 | -103.0965
20.482 | -103.0855
20.4857 | -103.0709
20.4895 | -103.0568
20.4932 | -103.0461
20.497 | -103.0356
20.5007 | -103.0238
20.5045 | -103.0098
20.5082 | -102.9985
20.512 | -102.9852
20.5157 | -102.9713
20.5195 | -102.961
20.5232 | -102.9489
20.527 | -102.933
20.5307 | -102.9191
20.5345 | -102.9072
20.5382 | -102.8969
20.542 | -102.8859
20.5457 | -102.8757
20.5495 | -102.8637
20.5532 | -102.8503
20.557 | -102.8414
20.5607 | -102.8289
20.5645 | -102.8158
20.5682 | -102.802
20.572 | -102.7872
20.5757 | -102.7726
20.5795 | -102.7579
20.5832 | -102.7419
20.587 | -102.7261
20.5907 | -102.7097
20.5945 | -102.6931
20.5982 | -102.6774
20.602 | -102.6613
20.6057 | -102.6454
20.6095 | -102.6282
20.6132 | -102.6122
20.617 | -102.5933
20.6207 | -102.5755
20.6245 | -102.5578
20.6282 | -102.5403
20.632 | -102.523
20.6357 | -102.5054
20.6395 | -102.4846
20.6432 | -102.4649
20.647 | -102.4477
20.6507 | -102.4301
20.6545 | -102.4122
20.6582 | -102.3939
20.662 | -102.3787
20.6657 | -102.364
20.6695 | -102.3481
20.6732 | -102.3323
20.677 | -102.3191
20.6807 | -102.3068
20.6845 | -102.298
20.6882 | -102.2895
20.692 | -102.2797
20.6957 | -102.2695
20.6995 | -102.2596
20.7032 | -102.2498
20.707 | -102.2411
20.7107 | -102.2352
20.7145 | -102.2314
20.7182 | -102.2273
20.722 | -102.2245
20.7257 | -102.2236
20.7295 | -102.2211
20.7332 | -102.2212
20.737 | -102.2218
20.7407 | -102.2264
20.7445 | -102.2324
20.7482 | -102.2389
20.752 | -102.2475
20.7557 | -102.2563
20.7595 | -102.2644
20.7632 | -102.2727
20.767 | -102.2831
20.7707 | -102.297
20.7745 | -102.3108
20.7782 | -102.3244
20.782 | -102.3352
20.7857 | -102.3459
20.7895 | -102.3573
20.7932 | -102.3696
20.797 | -102.3819
20.8008 | -102.3938
20.8045 | -102.4084
20.8083 | -102.4222
20.812 | -102.4359
20.8158 | -102.4489
20.8195 | -102.4628
20.8233 | -102.4764
20.827 | -102.4893
20.8308 | -102.4967
20.8345 | -102.503
20.8383 | -102.511
20.842 | -102.5178
20.8458 | -102.5274
20.8495 | -102.5341
20.8533 | -102.5389
20.857 | -102.5456
20.8608 | -102.5516
20.8645 | -102.5545
20.8683 | -102.5551
20.872 | -102.5524
20.8758 | -102.5493
20.8795 | -102.5467
20.8833 | -102.5414
20.887 | -102.5356
20.8908 | -102.5274
20.8945 | -102.5154
20.8983 | -102.5037
20.902 | -102.4897
20.9058 | -102.478
20.9095 | -102.4637
20.9133 | -102.449
20.917 | -102.4297
20.9208 | -102.4123
20.9245 | -102.3969
20.9283 | -102.3781
20.932 | -102.3605
20.9358 | -102.3448
20.9395 | -102.3293
20.9433 | -102.311
20.947 | -102.29
20.9508 | -102.2692
20.9545 | -102.2499
20.9583 | -102.2292
20.962 | -102.2096
20.9658 | -102.1923
20.9695 | -102.1748
20.9733 | -102.1548
20.977 | -102.138
20.9808 | -102.1189
20.9845 | -102.1006
20.9883 | -102.0838
20.992 | -102.0685
20.9958 | -102.0553
20.9995 | -102.0425
21.0033 | -102.0315
21.007 | -102.02
21.0108 | -102.0082
21.0145 | -101.9983
21.0183 | -101.9875
21.022 | -101.9775
21.0258 | -101.9702
21.0295 | -101.9622
21.0333 | -101.9543
21.037 | -101.9479
21.0408 | -101.9428
21.0445 | -101.9391
21.0483 | -101.9354
21.052 | -101.9333
21.0558 | -101.9325
21.0595 | -101.9351
21.0633 | -101.9382
21.067 | -101.9404
21.0708 | -101.9457
21.0745 | -101.9527
21.0783 | -101.9601
21.082 | -101.9693
21.0858 | -101.9808
21.0895 | -101.9934
21.0933 | -102.0075
21.097 | -102.0218
21.1008 | -102.0446
21.1045 | -102.0597
21.1083 | -102.0773
21.112 | -102.0947
21.1158 | -102.1133
21.1196 | -102.1333
21.1233 | -102.1524
21.1271 | -102.1725
21.1308 | -102.1948
21.1346 | -102.2192
21.1383 | -102.2423
21.1421 | -102.2653
21.1458 | -102.2852
21.1496 | -102.3056
21.1533 | -102.3272
21.1571 | -102.3487
21.1608 | -102.3705
21.1646 | -102.3937
21.1683 | -102.418
21.1721 | -102.4398
21.1758 | -102.4576
21.1796 | -102.4764
21.1833 | -102.4944
21.1871 | -102.511
21.1908 | -102.5263
21.1946 | -102.5413
21.1983 | -102.5524
21.2021 | -102.5646
21.2058 | -102.5767
21.2096 | -102.5875
21.2133 | -102.5947
21.2171 | -102.6032
21.2208 | -102.609
21.2246 | -102.6098
21.2283 | -102.6096
21.2321 | -102.6087
21.2358 | -102.6089
21.2396 | -102.6037
21.2433 | -102.5991
21.2471 | -102.5973
21.2508 | -102.5924
21.2546 | -102.5835
21.2583 | -102.5752
21.2621 | -102.5651
21.2658 | -102.5548
21.2696 | -102.5431
21.2733 | -102.531
21.2771 | -102.5126
21.2808 | -102.4981
21.2846 | -102.4859
21.2883 | -102.4722
21.2921 | -102.4561
21.2958 | -102.4409
21.2996 | -102.4256
21.3033 | -102.4088
21.3071 | -102.3883
21.3108 | -102.3665
21.3146 | -102.3441
21.3183 | -102.3267
21.3221 | -102.3106
21.3258 | -102.2956
21.3296 | -102.2808
21.3333 | -102.265
21.3371 | -102.2467
21.3408 | -102.2257
21.3446 | -102.2072
21.3483 | -102.1893
21.3521 | -102.1701
21.3558 | -102.1518
21.3596 | -102.1315
21.3633 | -102.1107
21.3671 | -102.0918
21.3708 | -102.0713
21.3746 | -102.05
21.3783 | -102.0288
21.3821 | -102.0059
21.3858 | -101.9836
21.3896 | -101.9603
21.3933 | -101.9372
21.3971 | -101.9161
21.4008 | -101.8911
21.4046 | -101.8634
21.4083 | -101.8381
21.4121 | -101.81
21.4158 | -101.7857
21.4196 | -101.762
21.4233 | -101.7377
21.4271 | -101.7135
21.4308 | -101.6883
21.4346 | -101.6609
21.4383 | -101.6314
21.4421 | -101.604
21.4459 | -101.5772
21.4496 | -101.5502
21.4534 | -101.5241
21.4571 | -101.4965
21.4609 | -101.4672
21.4646 | -101.4379
21.4684 | -101.4088
21.4721 | -101.3788
21.4759 | -101.3501
21.4796 | -101.3237
21.4834 | -101.2918
21.4871 | -101.2608
21.4909 | -101.2299
21.4946 | -101.1985
21.4984 | -101.168
21.5021 | -101.1387
21.5059 | -101.1055
21.5096 | -101.0704
21.5134 | -101.0357
21.5171 | -101.003
21.5209 | -100.9707
21.5246 | -100.9369
21.5284 | -100.9027
21.5321 | -100.8703
21.5359 | -100.8347
21.5396 | -100.8007
21.5434 | -100.7683
21.5471 | -100.7364
21.5509 | -100.7046
21.5546 | -100.6749
21.5584 | -100.6458
21.5621 | -100.615
21.5659 | -100.5874
21.5696 | -100.5608
21.5734 | -100.5336
21.5771 | -100.5086
21.5809 | -100.488
21.5846 | -100.4671
21.5884 | -100.4454
21.5921 | -100.4259
21.5959 | -100.4085
21.5996 | -100.3893
21.6034 | -100.3723
21.6071 | -100.358
21.6109 | -100.3473
21.6146 | -100.3377
21.6184 | -100.3272
21.6221 | -100.3159
21.6259 | -100.3063
21.6296 | -100.2979
21.6334 | -100.2913
21.6371 | -100.2866
21.6409 | -100.2862
21.6446 | -100.2809
21.6484 | -100.2768
21.6521 | -100.2725
21.6559 | -100.2686
21.6596 | -100.2646
21.6634 | -100.2605
21.6671 | -100.2567
21.6709 | -100.2526
21.6746 | -100.2497
21.6784 | -100.2461
21.6821 | -100.2428
21.6859 | -100.238
21.6896 | -100.2325
21.6934 | -100.2259
21.6971 | -100.2216
21.7009 | -100.2178
21.7046 | -100.2116
21.7084 | -100.2086
21.7121 | -100.204
21.7159 | -100.1948
21.7196 | -100.1842
21.7234 | -100.1704
21.7271 | -100.1539
21.7309 | -100.1359
21.7346 | -100.1181
21.7384 | -100.0987
21.7421 | -100.0829
21.7459 | -100.0667
21.7496 | -100.0478
21.7534 | -100.0281
21.7571 | -100.006
21.7609 | -99.9809
21.7647 | -99.9552
21.7684 | -99.9308
21.7722 | -99.9063
21.7759 | -99.8796
21.7797 | -99.8506
21.7834 | -99.8193
21.7872 | -99.7876
21.7909 | -99.7556
21.7947 | -99.723
21.7984 | -99.6896
21.8022 | -99.6558
21.8059 | -99.6227
21.8097 | -99.5917
21.8134 | -99.5618
21.8172 | -99.5318
21.8209 | -99.5003
21.8247 | -99.4698
21.8284 | -99.4412
21.8322 | -99.4129
21.8359 | -99.3828
21.8397 | -99.3539
21.8434 | -99.3294
21.8472 | -99.305
21.8509 | -99.2813
21.8547 | -99.2596
21.8584 | -99.2392
21.8622 | -99.2183
21.8659 | -99.1995
21.8697 | -99.182
21.8734 | -99.1662
21.8772 | -99.151
21.8809 | -99.1377
21.8847 | -99.1246
21.8884 | -99.1124
21.8922 | -99.1005
21.8959 | -99.091
21.8997 | -99.0816
21.9034 | -99.0754
21.9072 | -99.0731
21.9109 | -99.0704
21.9147 | -99.0661
21.9184 | -99.0619
21.9222 | -99.0607
21.9259 | -99.061
21.9297 | -99.062
21.9334 | -99.0617
21.9372 | -99.0625
21.9409 | -99.0657
21.9447 | -99.0701
21.9484 | -99.0746
21.9522 | -99.0784
21.9559 | -99.083
21.9597 | -99.0888
21.9634 | -99.0954
21.9672 | -99.1021
21.9709 | -99.1076
21.9747 | -99.1129
21.9784 | -99.1188
21.9822 | -99.1253
21.9859 | -99.1324
21.9897 | -99.1401
21.9934 | -99.1458
21.9972 | -99.1549
22.0009 | -99.1631
22.0047 | -99.1719
22.0084 | -99.1796
22.0122 | -99.1859
22.0159 | -99.1928
22.0197 | -99.1972
22.0234 | -99.2028
22.0272 | -99.2112
22.0309 | -99.2213
22.0347 | -99.2294
22.0384 | -99.2377
22.0422 | -99.2461
22.0459 | -99.2547
22.0497 | -99.2634
22.0534 | -99.2674
22.0572 | -99.2715
22.0609 | -99.2727
22.0647 | -99.2753
22.0684 | -99.2802
22.0722 | -99.2874
22.0759 | -99.2945
22.0797 | -99.2992
22.0835 | -99.3041
22.0872 | -99.3106
22.091 | -99.319
22.0947 | -99.3279
22.0985 | -99.3358
22.1022 | -99.3449
22.106 | -99.3548
22.1097 | -99.3654
22.1135 | -99.3761
22.1172 | -99.3855
22.121 | -99.393
22.1247 | -99.4023
22.1285 | -99.4127
22.1322 | -99.4242
22.136 | -99.435
22.1397 | -99.4441
22.1435 | -99.4537
22.1472 | -99.4674
22.151 | -99.4818
22.1547 | -99.4951
22.1585 | -99.5093
22.1622 | -99.5217
22.166 | -99.5316
22.1697 | -99.5434
22.1735 | -99.5557
22.1772 | -99.5684
22.181 | -99.583
22.1847 | -99.5957
22.1885 | -99.6091
22.1922 | -99.6234
22.196 | -99.6376
22.1997 | -99.6515
22.2035 | -99.6622
22.2072 | -99.6756
22.211 | -99.6899
22.2147 | -99.704
22.2185 | -99.7152
22.2222 | -99.7256
22.226 | -99.7366
22.2297 | -99.7504
22.2335 | -99.763
22.2372 | -99.7732
22.241 | -99.7845
22.2447 | -99.7951
22.2485 | -99.8077
22.2522 | -99.8201
22.256 | -99.8292
22.2597 | -99.8374
22.2635 | -99.849
22.2672 | -99.8619
22.271 | -99.8744
22.2747 | -99.8855
22.2785 | -99.8938
22.2822 | -99.902
22.286 | -99.9157
22.2897 | -99.9305
22.2935 | -99.9437
22.2972 | -99.9559
22.301 | -99.966
22.3047 | -99.9776
22.3085 | -99.9923
22.3122 | -100.0042
22.316 | -100.0175
22.3197 | -100.0306
22.3235 | -100.0443
22.3272 | -100.0598
22.331 | -100.0763
22.3347 | -100.0951
22.3385 | -100.1156
22.3422 | -100.1333
22.346 | -100.1509
22.3497 | -100.1704
22.3535 | -100.1925
22.3572 | -100.2147
22.361 | -100.2352
22.3647 | -100.2561
22.3685 | -100.2795
22.3722 | -100.3025
22.376 | -100.3297
22.3797 | -100.354
22.3835 | -100.3754
22.3872 | -100.3973
22.391 | -100.421
22.3947 | -100.4458
22.3985 | -100.4692
22.4022 | -100.4931
22.406 | -100.5149
22.4098 | -100.5385
22.4135 | -100.5623
22.4173 | -100.5849
22.421 | -100.6066
22.4248 | -100.6268
22.4285 | -100.6449
22.4323 | -100.6632
22.436 | -100.683
22.4398 | -100.7028
22.4435 | -100.7223
22.4473 | -100.7429
22.451 | -100.7622
22.4548 | -100.7815
22.4585 | -100.8017
22.4623 | -100.8207
22.466 | -100.8367
22.4698 | -100.8536
22.4735 | -100.87
22.4773 | -100.8887
22.481 | -100.9102
22.4848 | -100.9293
22.4885 | -100.9474
22.4923 | -100.9633
22.496 | -100.9789
22.4998 | -100.9976
22.5035 | -101.0166
22.5073 | -101.0327
22.511 | -101.0473
22.5148 | -101.0616
22.5185 | -101.076
22.5223 | -101.0943
22.526 | -101.1093
22.5298 | -101.1175
22.5335 | -101.1273
22.5373 | -101.1394
22.541 | -101.1501
22.5448 | -101.1573
22.5485 | -101.1623
22.5523 | -101.1681
22.556 | -101.1734
22.5598 | -101.1781
22.5635 | -101.1814
22.5673 | -101.1819
22.571 | -101.1806
22.5748 | -101.1804
22.5785 | -101.1804
22.5823 | -101.177
22.586 | -101.1733
22.5898 | -101.1679
22.5935 | -101.1612
22.5973 | -101.1568
22.601 | -101.1519
22.6048 | -101.1425
22.6085 | -101.1329
22.6123 | -101.1233
22.616 | -101.113
22.6198 | -101.1025
22.6235 | -101.091
22.6273 | -101.0791
22.631 | -101.068
22.6348 | -101.0566
22.6385 | -101.0457
22.6423 | -101.0356
22.646 | -101.0237
22.6498 | -101.0109
22.6535 | -101.0017
22.6573 | -100.992
22.661 | -100.9854
22.6648 | -100.978
22.6685 | -100.9711
22.6723 | -100.9644
22.676 | -100.9597
22.6798 | -100.9573
22.6835 | -100.957
22.6873 | -100.9586
22.691 | -100.9613
22.6948 | -100.9633
22.6985 | -100.9675
22.7023 | -100.9749
22.706 | -100.981
22.7098 | -100.9875
22.7135 | -100.9943
22.7173 | -101.0042
22.7211 | -101.0142
22.7248 | -101.0324
22.7286 | -101.0447
22.7323 | -101.0577
22.7361 | -101.0722
22.7398 | -101.0891
22.7436 | -101.1065
22.7473 | -101.1237
22.7511 | -101.141
22.7548 | -101.1589
22.7586 | -101.1748
22.7623 | -101.1903
22.7661 | -101.2052
22.7698 | -101.2206
22.7736 | -101.2353
22.7773 | -101.2491
22.7811 | -101.2623
22.7848 | -101.2718
22.7886 | -101.2822
22.7923 | -101.2912
22.7961 | -101.2978
22.7998 | -101.3051
22.8036 | -101.3109
22.8073 | -101.3154
22.8111 | -101.3193
22.8148 | -101.3231
22.8186 | -101.3257
22.8223 | -101.3254
22.8261 | -101.3234
22.8298 | -101.3202
22.8336 | -101.3195
22.8373 | -101.3204
22.8411 | -101.319
22.8448 | -101.3146
22.8486 | -101.3107
22.8523 | -101.3096
22.8561 | -101.3063
22.8598 | -101.2983
22.8636 | -101.2896
22.8673 | -101.282
22.8711 | -101.2721
22.8748 | -101.2629
22.8786 | -101.2519
22.8823 | -101.2399
22.8861 | -101.2285
22.8898 | -101.2168
22.8936 | -101.2032
22.8973 | -101.1869
22.9011 | -101.1703
22.9048 | -101.156
22.9086 | -101.1416
22.9123 | -101.1258
22.9161 | -101.1094
22.9198 | -101.092
22.9236 | -101.0745
22.9273 | -101.0576
22.9311 | -101.0408
22.9348 | -101.0222
22.9386 | -101.0038
22.9423 | -100.9864
22.9461 | -100.9682
22.9498 | -100.9498
22.9536 | -100.9299
22.9573 | -100.9111
22.9611 | -100.8916
22.9648 | -100.8709
22.9686 | -100.8491
22.9723 | -100.8284
22.9761 | -100.8071
22.9798 | -100.7852
22.9836 | -100.7639
22.9873 | -100.7449
22.9911 | -100.7263
22.9948 | -100.7073
22.9986 | -100.686
23.0023 | -100.6656
23.0061 | -100.645
23.0098 | -100.6245
23.0136 | -100.6054
23.0173 | -100.5871
23.0211 | -100.5692
23.0248 | -100.5506
23.0286 | -100.5322
23.0323 | -100.5159
23.0361 | -100.5007
23.0399 | -100.4844
23.0436 | -100.4706
23.0474 | -100.4595
23.0511 | -100.4482
23.0549 | -100.4379
23.0586 | -100.4292
23.0624 | -100.4211
23.0661 | -100.4127
23.0699 | -100.4054
23.0736 | -100.4015
23.0774 | -100.4014
23.0811 | -100.4013
23.0849 | -100.4004
23.0886 | -100.3994
23.0924 | -100.3992
23.0961 | -100.4004
23.0999 | -100.404
23.1036 | -100.4089
23.1074 | -100.4151
23.1111 | -100.4248
23.1149 | -100.4331
23.1186 | -100.4412
23.1224 | -100.452
23.1261 | -100.4632
23.1299 | -100.4747
23.1336 | -100.4881
23.1374 | -100.5031
23.1411 | -100.5176
23.1449 | -100.5318
23.1486 | -100.5457
23.1524 | -100.5616
23.1561 | -100.5802
23.1599 | -100.5961
23.1636 | -100.6112
23.1674 | -100.6279
23.1711 | -100.6459
23.1749 | -100.6639
23.1786 | -100.683
23.1824 | -100.7024
23.1861 | -100.7231
23.1899 | -100.7456
23.1936 | -100.7693
23.1974 | -100.794
23.2011 | -100.8171
23.2049 | -100.8403
23.2086 | -100.8635
23.2124 | -100.8869
23.2161 | -100.9111
23.2199 | -100.9381
23.2236 | -100.9671
23.2274 | -100.9959
23.2311 | -101.0232
23.2349 | -101.0499
23.2386 | -101.0778
23.2424 | -101.1055
23.2461 | -101.1354
23.2499 | -101.1656
23.2536 | -101.1929
23.2574 | -101.219
23.2611 | -101.2473
23.2649 | -101.2697
23.2686 | -101.2989
23.2724 | -101.327
23.2761 | -101.3545
23.2799 | -101.3811
23.2836 | -101.4075
23.2874 | -101.4323
23.2911 | -101.4552
23.2949 | -101.4766
23.2986 | -101.5005
23.3024 | -101.5239
23.3061 | -101.5409
23.3099 | -101.5566
23.3136 | -101.5745
23.3174 | -101.5905
23.3211 | -101.6044
23.3249 | -101.6186
23.3286 | -101.6334
23.3324 | -101.6447
23.3361 | -101.6548
23.3399 | -101.6639
23.3436 | -101.6728
23.3474 | -101.6823
23.3511 | -101.6908
23.3549 | -101.6967
23.3587 | -101.7025
23.3624 | -101.7077
23.3662 | -101.7108
23.3699 | -101.7138
23.3737 | -101.7161
23.3774 | -101.7156
23.3812 | -101.7133
23.3849 | -101.7132
23.3887 | -101.7116
23.3924 | -101.7095
23.3962 | -101.7075
23.3999 | -101.7062
23.4037 | -101.7034
23.4074 | -101.7002
23.4112 | -101.6976
23.4149 | -101.6957
23.4187 | -101.6934
23.4224 | -101.6932
23.4262 | -101.6929
23.4299 | -101.6918
23.4337 | -101.6893
23.4374 | -101.6866
23.4412 | -101.6837
23.4449 | -101.6799
23.4487 | -101.6772
23.4524 | -101.6755
23.4562 | -101.6733
23.4599 | -101.6721
23.4637 | -101.6714
23.4674 | -101.6689
23.4712 | -101.6693
23.4749 | -101.6681
23.4787 | -101.6691
23.4824 | -101.6703
23.4862 | -101.671
23.4899 | -101.6734
23.4937 | -101.6763
23.4974 | -101.6786
23.5012 | -101.6808
23.5049 | -101.6826
23.5087 | -101.683
23.5124 | -101.6831
23.5162 | -101.6842
23.5199 | -101.684
23.5237 | -101.6846
23.5274 | -101.6855
23.5312 | -101.6868
23.5349 | -101.6885
23.5387 | -101.6892
23.5424 | -101.6891
23.5462 | -101.69
23.5499 | -101.6905
23.5537 | -101.6911
23.5574 | -101.6903
23.5612 | -101.6888
23.5649 | -101.6882
23.5687 | -101.6884
23.5724 | -101.6905
23.5762 | -101.69
23.5799 | -101.6884
23.5837 | -101.6891
23.5874 | -101.6914
23.5912 | -101.6948
23.5949 | -101.6992
23.5987 | -101.7009
23.6024 | -101.7008
23.6062 | -101.6994
23.6099 | -101.7
23.6137 | -101.7015
23.6174 | -101.7053
23.6212 | -101.7094
23.6249 | -101.7136
23.6287 | -101.7182
23.6324 | -101.7234
23.6362 | -101.7285
23.6399 | -101.7369
23.6437 | -101.7452
23.6474 | -101.7509
23.6512 | -101.7578
23.6549 | -101.7661
23.6587 | -101.7732
23.6624 | -101.782
23.6662 | -101.7929
23.6699 | -101.8061
23.6737 | -101.8194
23.6775 | -101.8296
23.6812 | -101.8395
23.685 | -101.8505
23.6887 | -101.866
23.6925 | -101.8805
23.6962 | -101.8917
23.7 | -101.9054
23.7037 | -101.9184
23.7075 | -101.9301
23.7112 | -101.9437
23.715 | -101.9585
23.7187 | -101.9726
23.7225 | -101.9846
23.7262 | -101.9951
23.73 | -102.0069
23.7337 | -102.0189
23.7375 | -102.0313
23.7412 | -102.0429
23.745 | -102.0536
23.7487 | -102.0648
23.7525 | -102.0748
23.7562 | -102.0825
23.76 | -102.0888
23.7637 | -102.096
23.7675 | -102.103
23.7712 | -102.1094
23.775 | -102.115
23.7787 | -102.1212
23.7825 | -102.1239
23.7862 | -102.1249
23.79 | -102.1265
23.7937 | -102.1251
23.7975 | -102.1221
23.8012 | -102.1219
23.805 | -102.1144
23.8087 | -102.1112
23.8125 | -102.1102
23.8162 | -102.1061
23.82 | -102.1
23.8237 | -102.0943
23.8275 | -102.0889
23.8312 | -102.0838
23.835 | -102.0792
23.8387 | -102.0738
23.8425 | -102.0677
23.8462 | -102.0622
23.85 | -102.0573
23.8537 | -102.051
23.8575 | -102.0448
23.8612 | -102.0381
23.865 | -102.0325
23.8687 | -102.0261
23.8725 | -102.0211
23.8762 | -102.016
23.88 | -102.0117
23.8837 | -102.0067
23.8875 | -102.0008
23.8912 | -101.9964
23.895 | -101.9957
23.8987 | -101.9951
23.9025 | -101.9942
23.9062 | -101.9925
23.91 | -101.9883
23.9137 | -101.9861
23.9175 | -101.987
23.9212 | -101.9884
23.925 | -101.9879
23.9287 | -101.9859
23.9325 | -101.9829
23.9362 | -101.9839
23.94 | -101.9869
23.9437 | -101.9919
23.9475 | -101.9956
23.9512 | -101.9982
23.955 | -102.0003
23.9587 | -102.0046
23.9625 | -102.0104
23.9662 | -102.0142
23.97 | -102.0181
23.9737 | -102.0218
23.9775 | -102.0266
23.9812 | -102.0325
23.985 | -102.0351
23.9887 | -102.0362
23.9925 | -102.0363
23.9962 | -102.038
24 | -102.0399
24.0038 | -102.0401
24.0075 | -102.0375
24.0113 | -102.0333
24.015 | -102.0291
24.0188 | -102.0226
24.0225 | -102.0147
24.0263 | -102.0065
24.03 | -101.997
24.0338 | -101.9875
24.0375 | -101.9765
24.0413 | -101.967
24.045 | -101.9561
24.0488 | -101.9431
24.0525 | -101.9312
24.0563 | -101.9152
24.06 | -101.9009
24.0638 | -101.8875
24.0675 | -101.8725
24.0713 | -101.8548
24.075 | -101.8374
24.0788 | -101.8215
24.0825 | -101.804
24.0863 | -101.7855
24.09 | -101.7687
24.0938 | -101.7511
24.0975 | -101.7337
24.1013 | -101.7168
24.105 | -101.7014
24.1088 | -101.6858
24.1125 | -101.6701
24.1163 | -101.6552
24.12 | -101.642
24.1238 | -101.63
24.1275 | -101.6187
24.1313 | -101.6093
24.135 | -101.5997
24.1388 | -101.5894
24.1425 | -101.5807
24.1463 | -101.5734
24.15 | -101.5682
24.1538 | -101.5617
24.1575 | -101.5561
24.1613 | -101.5539
24.165 | -101.552
24.1688 | -101.5521
24.1725 | -101.553
24.1763 | -101.5561
24.18 | -101.5589
24.1838 | -101.5623
24.1875 | -101.5689
24.1913 | -101.5757
24.195 | -101.5838
24.1988 | -101.5918
24.2025 | -101.5998
24.2063 | -101.6078
24.21 | -101.6158
24.2138 | -101.6235
24.2175 | -101.6325
24.2213 | -101.6428
24.225 | -101.6545
24.2288 | -101.6679
24.2325 | -101.6789
24.2363 | -101.6889
24.24 | -101.6979
24.2438 | -101.7094
24.2475 | -101.7222
24.2513 | -101.7344
24.255 | -101.7444
24.2588 | -101.7509
24.2625 | -101.7576
24.2663 | -101.7657
24.27 | -101.7731
24.2738 | -101.7789
24.2775 | -101.7777
24.2813 | -101.7735
24.285 | -101.7617
24.2888 | -101.752
24.2925 | -101.7463
24.2963 | -101.7391
24.3 | -101.7296
24.3038 | -101.7189
24.3075 | -101.7094
24.3113 | -101.6969
24.315 | -101.6818
24.3188 | -101.6664
24.3225 | -101.65
24.3263 | -101.6342
24.3301 | -101.6176
24.3338 | -101.6009
24.3376 | -101.5822
24.3413 | -101.5616
24.3451 | -101.541
24.3488 | -101.5139
24.3526 | -101.4941
24.3563 | -101.4748
24.3601 | -101.4543
24.3638 | -101.434
24.3676 | -101.4129
24.3713 | -101.3905
24.3751 | -101.3687
24.3788 | -101.3487
24.3826 | -101.328
24.3863 | -101.3072
24.3901 | -101.2872
24.3938 | -101.2683
24.3976 | -101.2497
24.4013 | -101.2344
24.4051 | -101.2196
24.4088 | -101.2058
24.4126 | -101.1933
24.4163 | -101.1848
24.4201 | -101.1769
24.4238 | -101.1685
24.4276 | -101.1656
24.4313 | -101.1652
24.4351 | -101.1647
24.4388 | -101.166
24.4426 | -101.1688
24.4463 | -101.1732
24.4501 | -101.1811
24.4538 | -101.1914
24.4576 | -101.2032
24.4613 | -101.2178
24.4651 | -101.236
24.4688 | -101.2542
24.4726 | -101.2734
24.4763 | -101.2949
24.4801 | -101.3198
24.4838 | -101.346
24.4876 | -101.3728
24.4913 | -101.4008
24.4951 | -101.4306
24.4988 | -101.4626
24.5026 | -101.4957
24.5063 | -101.5279
24.5101 | -101.5623
24.5138 | -101.5958
24.5176 | -101.6297
24.5213 | -101.6635
24.5251 | -101.6986
24.5288 | -101.7343
24.5326 | -101.7698
24.5363 | -101.8058
24.5401 | -101.8407
24.5438 | -101.8753
24.5476 | -101.9106
24.5513 | -101.9453
24.5551 | -101.9809
24.5588 | -102.0138
24.5626 | -102.0453
24.5663 | -102.0737
24.5701 | -102.0999
24.5738 | -102.1227
24.5776 | -102.1452
24.5813 | -102.164
24.5851 | -102.1747
24.5888 | -102.1719
24.5926 | -102.1476
24.5963 | -102.1266
24.6001 | -102.1161
24.6038 | -102.1097
24.6076 | -102.1029
24.6113 | -102.094
24.6151 | -102.0817
24.6188 | -102.0665
24.6226 | -102.053
24.6263 | -102.0358
24.6301 | -102.0174
24.6338 | -101.9934
24.6376 | -101.962
24.6413 | -101.9264
24.6451 | -101.8866
24.6489 | -101.8435
24.6526 | -101.7988
24.6564 | -101.7505
24.6601 | -101.6987
24.6639 | -101.6421
24.6676 | -101.5844
24.6714 | -101.5278
24.6751 | -101.471
24.6789 | -101.4136
24.6826 | -101.3534
24.6864 | -101.2928
24.6901 | -101.2329
24.6939 | -101.1746
24.6976 | -101.1163
24.7014 | -101.0587
24.7051 | -101.0019
24.7089 | -100.9484
24.7126 | -100.896
24.7164 | -100.8446
24.7201 | -100.7959
24.7239 | -100.7505
24.7276 | -100.7066
24.7314 | -100.6655
24.7351 | -100.6271
24.7389 | -100.593
24.7426 | -100.5623
24.7464 | -100.5359
24.7501 | -100.5137
24.7539 | -100.4955
24.7576 | -100.4814
24.7614 | -100.4722
24.7651 | -100.4662
24.7689 | -100.4647
24.7726 | -100.4689
24.7764 | -100.4762
24.7801 | -100.4891
24.7839 | -100.5086
24.7876 | -100.5322
24.7914 | -100.5592
24.7951 | -100.5906
24.7989 | -100.6254
24.8026 | -100.6642
24.8064 | -100.7073
24.8101 | -100.7547
24.8139 | -100.8045
24.8176 | -100.8582
24.8214 | -100.9156
24.8251 | -100.9756
24.8289 | -101.0372
24.8326 | -101.1028
24.8364 | -101.1713
24.8401 | -101.2419
24.8439 | -101.3126
24.8476 | -101.3854
24.8514 | -101.4617
24.8551 | -101.5382
24.8589 | -101.615
24.8626 | -101.6919
24.8664 | -101.7679
24.8701 | -101.8426
24.8739 | -101.9184
24.8776 | -101.994
24.8814 | -102.0668
24.8851 | -102.1373
24.8889 | -102.1811
24.8926 | -102.2471
24.8964 | -102.3107
24.9001 | -102.3716
24.9039 | -102.4307
24.9076 | -102.4867
24.9114 | -102.5403
24.9151 | -102.5919
24.9189 | -102.6414
24.9226 | -102.6868
24.9264 | -102.7325
24.9301 | -102.7758
24.9339 | -102.8164
24.9376 | -102.861
24.9414 | -102.9108
24.9451 | -102.9738
24.9489 | -103.0228
24.9526 | -103.0502
24.9564 | -103.0665
24.9601 | -103.0776
24.9639 | -103.0854
24.9677 | -103.0932
24.9714 | -103.1002
24.9752 | -103.1062
24.9789 | -103.1113
24.9827 | -103.1194
24.9864 | -103.1222
24.9902 | -103.1237
24.9939 | -103.1293
24.9977 | -103.1348
25.0014 | -103.1413
25.0052 | -103.1479
25.0089 | -103.1541
25.0127 | -103.1655
25.0164 | -103.1765
25.0202 | -103.1865
25.0239 | -103.1943
25.0277 | -103.2012
25.0314 | -103.2083
25.0352 | -103.2166
25.0389 | -103.2243
25.0427 | -103.231
25.0464 | -103.2391
25.0502 | -103.2475
25.0539 | -103.2549
25.0577 | -103.2628
25.0614 | -103.27
25.0652 | -103.279
25.0689 | -103.2875
25.0727 | -103.2958
25.0764 | -103.3034
25.0802 | -103.3119
25.0839 | -103.322
25.0877 | -103.3323
25.0914 | -103.3406
25.0952 | -103.3483
25.0989 | -103.3553
25.1027 | -103.3609
25.1064 | -103.3677
25.1102 | -103.3738
25.1139 | -103.3783
25.1177 | -103.3812
25.1214 | -103.3843
25.1252 | -103.3862
25.1289 | -103.3884
25.1327 | -103.3865
25.1364 | -103.3844
25.1402 | -103.3835
25.1439 | -103.3838
25.1477 | -103.3834
25.1514 | -103.3828
25.1552 | -103.3809
25.1589 | -103.3795
25.1627 | -103.3765
25.1664 | -103.3727
25.1702 | -103.3692
25.1739 | -103.3675
25.1777 | -103.3661
25.1814 | -103.3643
25.1852 | -103.3605
25.1889 | -103.3558
25.1927 | -103.3522
25.1964 | -103.3497
25.2002 | -103.3444
25.2039 | -103.3384
25.2077 | -103.334
25.2114 | -103.3288
25.2152 | -103.3237
25.2189 | -103.3204
25.2227 | -103.3188
25.2264 | -103.3158
25.2302 | -103.3121
25.2339 | -103.3094
25.2377 | -103.3083
25.2414 | -103.3059
25.2452 | -103.3033
25.2489 | -103.3011
25.2527 | -103.2991
25.2564 | -103.2976
25.2602 | -103.2969
25.2639 | -103.2952
25.2677 | -103.2948
25.2714 | -103.2949
25.2752 | -103.2963
25.2789 | -103.2986
25.2827 | -103.2987
25.2865 | -103.2986
25.2902 | -103.3014
25.294 | -103.3036
25.2977 | -103.3036
25.3015 | -103.3043
25.3052 | -103.3057
25.309 | -103.3072
25.3127 | -103.3094
25.3165 | -103.3106
25.3202 | -103.3132
25.324 | -103.3169
25.3277 | -103.3231
25.3315 | -103.326
25.3352 | -103.3291
25.339 | -103.333
25.3427 | -103.3397
25.3465 | -103.3448
25.3502 | -103.3468
25.354 | -103.3502
25.3577 | -103.3539
25.3615 | -103.3571
25.3652 | -103.3626
25.369 | -103.3654
25.3727 | -103.3698
25.3765 | -103.372
25.3802 | -103.3734
25.384 | -103.3748
25.3877 | -103.3769
25.3915 | -103.3812
25.3952 | -103.3853
25.399 | -103.3883
25.4027 | -103.3879
25.4065 | -103.3864
25.4102 | -103.3849
25.414 | -103.3846
25.4177 | -103.3838
25.4215 | -103.38
25.4252 | -103.3773
25.429 | -103.3715
25.4327 | -103.3671
25.4365 | -103.363
25.4402 | -103.3579
25.444 | -103.3534
25.4477 | -103.3497
25.4515 | -103.3467
25.4552 | -103.3435
25.459 | -103.3386
25.4627 | -103.3337
25.4665 | -103.3284
25.4702 | -103.3218
25.474 | -103.3165
25.4777 | -103.3125
25.4815 | -103.3075
25.4852 | -103.3019
25.489 | -103.2949
25.4927 | -103.291
25.4965 | -103.2849
25.5002 | -103.2788
25.504 | -103.2738
25.5077 | -103.2689
25.5115 | -103.2628
25.5152 | -103.2591
25.519 | -103.2567
25.5227 | -103.2535
25.5265 | -103.2512
25.5302 | -103.2492
25.534 | -103.2461
25.5377 | -103.2438
25.5415 | -103.242
25.5452 | -103.2398
25.549 | -103.2385
25.5527 | -103.2359
25.5565 | -103.234
25.5602 | -103.235
25.564 | -103.2357
25.5677 | -103.2351
25.5715 | -103.2365
25.5752 | -103.2398
25.579 | -103.2425
25.5827 | -103.244
25.5865 | -103.2446
25.5902 | -103.2468
25.594 | -103.2485
25.5978 | -103.2512
25.6015 | -103.255
25.6053 | -103.2582
25.609 | -103.2625
25.6128 | -103.2694
25.6165 | -103.2749
25.6203 | -103.279
25.624 | -103.2848
25.6278 | -103.2903
25.6315 | -103.2949
25.6353 | -103.2996
25.639 | -103.3037
25.6428 | -103.3083
25.6465 | -103.3126
25.6503 | -103.3154
25.654 | -103.3198
25.6578 | -103.3253
25.6615 | -103.33
25.6653 | -103.3334
25.669 | -103.3358
25.6728 | -103.3379
25.6765 | -103.3411
25.6803 | -103.3432
25.684 | -103.3454
25.6878 | -103.3451
25.6915 | -103.3472
25.6953 | -103.3508
25.699 | -103.3516
25.7028 | -103.3502
25.7065 | -103.3502
25.7103 | -103.35
25.714 | -103.3501
25.7178 | -103.3476
25.7215 | -103.3433
25.7253 | -103.3377
25.729 | -103.3323
25.7328 | -103.3281
25.7365 | -103.3243
25.7403 | -103.3192
25.744 | -103.3137
25.7478 | -103.3067
25.7515 | -103.2993
25.7553 | -103.2905
25.759 | -103.2829
25.7628 | -103.2761
25.7665 | -103.2698
25.7703 | -103.2623
25.774 | -103.2532
25.7778 | -103.2442
25.7815 | -103.2347
25.7853 | -103.2249
25.789 | -103.2174
25.7928 | -103.2098
25.7965 | -103.2022
25.8003 | -103.1942
25.804 | -103.186
25.8078 | -103.1766
25.8115 | -103.1676
25.8153 | -103.1594
25.819 | -103.1518
25.8228 | -103.1439
25.8265 | -103.1373
25.8303 | -103.1312
25.834 | -103.127
25.8378 | -103.1232
25.8415 | -103.1196
25.8453 | -103.1169
25.849 | -103.1137
25.8528 | -103.1103
25.8565 | -103.1071
25.8603 | -103.1044
25.864 | -103.102
25.8678 | -103.0999
25.8715 | -103.0975
25.8753 | -103.0968
25.879 | -103.0962
25.8828 | -103.0958
25.8865 | -103.0962
25.8903 | -103.0962
25.894 | -103.0953
25.8978 | -103.0965
25.9015 | -103.0971
25.9053 | -103.0989
25.909 | -103.1025
25.9128 | -103.1055
25.9165 | -103.1099
25.9203 | -103.1125
25.9241 | -103.1167
25.9278 | -103.1202
25.9316 | -103.1256
25.9353 | -103.1303
25.9391 | -103.1341
25.9428 | -103.1389
25.9466 | -103.144
25.9503 | -103.1498
25.9541 | -103.1548
25.9578 | -103.162
25.9616 | -103.1688
25.9653 | -103.1766
25.9691 | -103.1878
25.9728 | -103.1966
25.9766 | -103.2045
25.9803 | -103.2097
25.9841 | -103.2161
25.9878 | -103.2246
25.9916 | -103.2317
25.9953 | -103.24
25.9991 | -103.2499
26.0028 | -103.2579
26.0066 | -103.2671
26.0103 | -103.2767
26.0141 | -103.2873
26.0178 | -103.2968
26.0216 | -103.3064
26.0253 | -103.3152
26.0291 | -103.324
26.0328 | -103.3339
26.0366 | -103.3437
26.0403 | -103.3524
26.0441 | -103.3592
26.0478 | -103.3647
26.0516 | -103.3691
26.0553 | -103.3755
26.0591 | -103.3814
26.0628 | -103.3855
26.0666 | -103.391
26.0703 | -103.3977
26.0741 | -103.4023
26.0778 | -103.4044
26.0816 | -103.4068
26.0853 | -103.4105
26.0891 | -103.4145
26.0928 | -103.4178
26.0966 | -103.4206
26.1003 | -103.4209
26.1041 | -103.4223
26.1078 | -103.4238
26.1116 | -103.4251
26.1153 | -103.4264
26.1191 | -103.4273
26.1228 | -103.4286
26.1266 | -103.4291
26.1303 | -103.428
26.1341 | -103.4266
26.1378 | -103.4267
26.1416 | -103.4265
26.1453 | -103.4259
26.1491 | -103.4264
26.1528 | -103.4266
26.1566 | -103.4244
26.1603 | -103.4227
26.1641 | -103.4217
26.1678 | -103.4208
26.1716 | -103.4207
26.1753 | -103.4219
26.1791 | -103.4219
26.1828 | -103.4213
26.1866 | -103.4189
26.1903 | -103.4184
26.1941 | -103.4159
26.1978 | -103.4142
26.2016 | -103.4126
26.2053 | -103.4113
26.2091 | -103.4089
26.2128 | -103.4081
26.2166 | -103.4095
26.2203 | -103.4097
26.2241 | -103.409
26.2278 | -103.4082
26.2316 | -103.4066
26.2353 | -103.403
26.2391 | -103.4001
26.2429 | -103.398
26.2466 | -103.3959
26.2504 | -103.3935
26.2541 | -103.3912
26.2579 | -103.3884
26.2616 | -103.3878
26.2654 | -103.3842
26.2691 | -103.379
26.2729 | -103.3756
26.2766 | -103.3713
26.2804 | -103.3671
26.2841 | -103.3626
26.2879 | -103.3592
26.2916 | -103.3522
26.2954 | -103.3458
26.2991 | -103.3396
26.3029 | -103.3322
26.3066 | -103.3263
26.3104 | -103.3195
26.3141 | -103.3132
26.3179 | -103.3068
26.3216 | -103.2996
26.3254 | -103.2902
26.3291 | -103.2805
26.3329 | -103.27
26.3366 | -103.2565
26.3404 | -103.2448
26.3441 | -103.2348
26.3479 | -103.2239
26.3516 | -103.2121
26.3554 | -103.2007
26.3591 | -103.1878
26.3629 | -103.1758
26.3666 | -103.1633
26.3704 | -103.1507
26.3741 | -103.1381
26.3779 | -103.1254
26.3816 | -103.1111
26.3854 | -103.0974
26.3891 | -103.085
26.3929 | -103.0711
26.3966 | -103.0579
26.4004 | -103.0449
26.4041 | -103.0307
26.4079 | -103.017
26.4116 | -103.0048
26.4154 | -102.9921
26.4191 | -102.9808
26.4229 | -102.9705
26.4266 | -102.959
26.4304 | -102.9467
26.4341 | -102.9352
26.4379 | -102.923
26.4416 | -102.9118
26.4454 | -102.9025
26.4491 | -102.8951
26.4529 | -102.8862
26.4566 | -102.8768
26.4604 | -102.8695
26.4641 | -102.8632
26.4679 | -102.8545
26.4716 | -102.8473
26.4754 | -102.8392
26.4791 | -102.8313
26.4829 | -102.8258
26.4866 | -102.8198
26.4904 | -102.8134
26.4941 | -102.8084
26.4979 | -102.8054
26.5016 | -102.8032
26.5054 | -102.8009
26.5091 | -102.7989
26.5129 | -102.7983
26.5166 | -102.7953
26.5204 | -102.7917
26.5241 | -102.7872
26.5279 | -102.7825
26.5316 | -102.7781
26.5354 | -102.775
26.5391 | -102.7715
26.5429 | -102.7689
26.5466 | -102.7663
26.5504 | -102.7631
26.5541 | -102.7606
26.5579 | -102.7583
26.5617 | -102.7569
26.5654 | -102.7548
26.5692 | -102.7526
26.5729 | -102.7487
26.5767 | -102.7443
26.5804 | -102.7372
26.5842 | -102.7311
26.5879 | -102.7251
26.5917 | -102.7195
26.5954 | -102.7153
26.5992 | -102.7095
26.6029 | -102.7026
26.6067 | -102.6965
26.6104 | -102.6904
26.6142 | -102.6843
26.6179 | -102.6767
26.6217 | -102.6683
26.6254 | -102.6589
26.6292 | -102.6509
26.6329 | -102.6438
26.6367 | -102.6344
26.6404 | -102.6237
26.6442 | -102.6143
26.6479 | -102.6027
26.6517 | -102.593
26.6554 | -102.5842
26.6592 | -102.5748
26.6629 | -102.5643
26.6667 | -102.555
26.6704 | -102.5471
26.6742 | -102.538
26.6779 | -102.529
26.6817 | -102.5199
26.6854 | -102.5129
26.6892 | -102.5049
26.6929 | -102.4979
26.6967 | -102.4924
26.7004 | -102.4871
26.7042 | -102.4821
26.7079 | -102.4786
26.7117 | -102.4761
26.7154 | -102.4734
26.7192 | -102.4694
26.7229 | -102.4669
26.7267 | -102.4674
26.7304 | -102.4672
26.7342 | -102.4671
26.7379 | -102.4671
26.7417 | -102.4675
26.7454 | -102.4689
26.7492 | -102.4714
26.7529 | -102.474
26.7567 | -102.4779
26.7604 | -102.4817
26.7642 | -102.4865
26.7679 | -102.4932
26.7717 | -102.5008
26.7754 | -102.5081
26.7792 | -102.5158
26.7829 | -102.5236
26.7867 | -102.5324
26.7904 | -102.5392
26.7942 | -102.5458
26.7979 | -102.5543
26.8017 | -102.5633
26.8054 | -102.5721
26.8092 | -102.5818
26.8129 | -102.5912
26.8167 | -102.6002
26.8204 | -102.6091
26.8242 | -102.6182
26.8279 | -102.6268
26.8317 | -102.6354
26.8354 | -102.6421
26.8392 | -102.6488
26.8429 | -102.6578
26.8467 | -102.665
26.8504 | -102.6717
26.8542 | -102.6774
26.8579 | -102.6812
26.8617 | -102.6871
26.8654 | -102.6931
26.8692 | -102.6971
26.8729 | -102.6979
26.8767 | -102.6997
26.8804 | -102.7005
26.8842 | -102.7019
26.888 | -102.7028
26.8917 | -102.7024
26.8955 | -102.7032
26.8992 | -102.7033
26.903 | -102.7037
26.9067 | -102.7022
26.9105 | -102.7009
26.9142 | -102.6988
26.918 | -102.6975
26.9217 | -102.695
26.9255 | -102.6906
26.9292 | -102.6875
26.933 | -102.6832
26.9367 | -102.6775
26.9405 | -102.6725
26.9442 | -102.6679
26.948 | -102.6622
26.9517 | -102.6563
26.9555 | -102.6515
26.9592 | -102.6487
26.963 | -102.6439
26.9667 | -102.6382
26.9705 | -102.6326
26.9742 | -102.6275
26.978 | -102.6235
26.9817 | -102.6201
26.9855 | -102.6163
26.9892 | -102.6139
26.993 | -102.6107
26.9967 | -102.6073
27.0005 | -102.6055
27.0042 | -102.6045
27.008 | -102.6017
27.0117 | -102.5999
27.0155 | -102.5996
27.0192 | -102.6011
27.023 | -102.6024
27.0267 | -102.6031
27.0305 | -102.6038
27.0342 | -102.6048
27.038 | -102.6062
27.0417 | -102.6081
27.0455 | -102.6109
27.0492 | -102.612
27.053 | -102.6195
27.0567 | -102.6222
27.0605 | -102.6258
27.0642 | -102.6293
27.068 | -102.6327
27.0717 | -102.6377
27.0755 | -102.6422
27.0792 | -102.6466
27.083 | -102.6514
27.0867 | -102.6568
27.0905 | -102.6631
27.0942 | -102.6696
27.098 | -102.6748
27.1017 | -102.6807
27.1055 | -102.687
27.1092 | -102.6925
27.113 | -102.6976
27.1167 | -102.7031
27.1205 | -102.7099
27.1242 | -102.7162
27.128 | -102.7235
27.1317 | -102.7316
27.1355 | -102.7404
27.1392 | -102.7491
27.143 | -102.7573
27.1467 | -102.7652
27.1505 | -102.7717
27.1542 | -102.7777
27.158 | -102.7834
27.1617 | -102.7897
27.1655 | -102.7962
27.1692 | -102.8013
27.173 | -102.81
27.1767 | -102.8182
27.1805 | -102.8243
27.1842 | -102.8305
27.188 | -102.8357
27.1917 | -102.8393
27.1955 | -102.8441
27.1992 | -102.849
27.203 | -102.8541
27.2068 | -102.8587
27.2105 | -102.8635
27.2143 | -102.8672
27.218 | -102.8711
27.2218 | -102.8748
27.2255 | -102.878
27.2293 | -102.8818
27.233 | -102.8847
27.2368 | -102.8882
27.2405 | -102.8912
27.2443 | -102.8914
27.248 | -102.8937
27.2518 | -102.896
27.2555 | -102.8999
27.2593 | -102.9025
27.263 | -102.9064
27.2668 | -102.9083
27.2705 | -102.9107
27.2743 | -102.9142
27.278 | -102.9181
27.2818 | -102.9208
27.2855 | -102.9217
27.2893 | -102.9246
27.293 | -102.9264
27.2968 | -102.9269
27.3005 | -102.9286
27.3043 | -102.9311
27.308 | -102.9342
27.3118 | -102.937
27.3155 | -102.9408
27.3193 | -102.9444
27.323 | -102.9482
27.3268 | -102.9525
27.3305 | -102.9578
27.3343 | -102.962
27.338 | -102.9646
27.3418 | -102.9678
27.3455 | -102.9712
27.3493 | -102.9758
27.353 | -102.9802
27.3568 | -102.9844
27.3605 | -102.9897
27.3643 | -102.9947
27.368 | -102.9992
27.3718 | -103.0043
27.3755 | -103.01
27.3793 | -103.016
27.383 | -103.0212
27.3868 | -103.0262
27.3905 | -103.0319
27.3943 | -103.0367
27.398 | -103.0412
27.4018 | -103.0466
27.4055 | -103.0516
27.4093 | -103.0578
27.413 | -103.0625
27.4168 | -103.0674
27.4205 | -103.0708
27.4243 | -103.0755
27.428 | -103.0795
27.4318 | -103.0846
27.4355 | -103.0881
27.4393 | -103.0913
27.443 | -103.0953
27.4468 | -103.0994
27.4505 | -103.1026
27.4543 | -103.1056
27.458 | -103.1079
27.4618 | -103.1113
27.4655 | -103.1142
27.4693 | -103.1197
27.473 | -103.1231
27.4768 | -103.1241
27.4805 | -103.1256
27.4843 | -103.1285
27.488 | -103.1314
27.4918 | -103.1332
27.4955 | -103.1356
27.4993 | -103.1365
27.503 | -103.1362
27.5068 | -103.1345
27.5105 | -103.1328
27.5143 | -103.1279
27.518 | -103.1216
27.5218 | -103.1154
27.5256 | -103.1106
27.5293 | -103.1059
27.5331 | -103.0991
27.5368 | -103.0937
27.5406 | -103.0878
27.5443 | -103.0819
27.5481 | -103.0755
27.5518 | -103.0687
27.5556 | -103.0608
27.5593 | -103.0512
27.5631 | -103.0397
27.5668 | -103.027
27.5706 | -103.0157
27.5743 | -103.0072
27.5781 | -102.9973
27.5818 | -102.9864
27.5856 | -102.9753
27.5893 | -102.9619
27.5931 | -102.9503
27.5968 | -102.9395
27.6006 | -102.9293
27.6043 | -102.9198
27.6081 | -102.91
27.6118 | -102.9003
27.6156 | -102.8903
27.6193 | -102.8795
27.6231 | -102.8697
27.6268 | -102.8596
27.6306 | -102.8503
27.6343 | -102.8413
27.6381 | -102.8305
27.6418 | -102.8212
27.6456 | -102.8108
27.6493 | -102.8027
27.6531 | -102.7953
27.6568 | -102.7895
27.6606 | -102.7841
27.6643 | -102.7792
27.6681 | -102.7737
27.6718 | -102.7696
27.6756 | -102.7679
27.6793 | -102.7664
27.6831 | -102.7648
27.6868 | -102.7633
27.6906 | -102.7623
27.6943 | -102.7616
27.6981 | -102.7618
27.7018 | -102.7615
27.7056 | -102.7618
27.7093 | -102.7635
27.7131 | -102.7656
27.7168 | -102.7671
27.7206 | -102.7695
27.7243 | -102.773
27.7281 | -102.7764
27.7318 | -102.7798
27.7356 | -102.7851
27.7393 | -102.7892
27.7431 | -102.7928
27.7468 | -102.7976
27.7506 | -102.8025
27.7543 | -102.808
27.7581 | -102.8133
27.7618 | -102.8175
27.7656 | -102.8223
27.7693 | -102.828
27.7731 | -102.8342
27.7768 | -102.8402
27.7806 | -102.8462
27.7843 | -102.854
27.7881 | -102.8618
27.7918 | -102.8689
27.7956 | -102.8762
27.7993 | -102.883
27.8031 | -102.8902
27.8068 | -102.8968
27.8106 | -102.9058
27.8143 | -102.9147
27.8181 | -102.9235
27.8218 | -102.9296
27.8256 | -102.9357
27.8293 | -102.9435
27.8331 | -102.9509
27.8368 | -102.9584
27.8406 | -102.9665
27.8444 | -102.9744
27.8481 | -102.9819
27.8519 | -102.987
27.8556 | -102.9938
27.8594 | -102.9996
27.8631 | -103.0057
27.8669 | -103.0131
27.8706 | -103.0208
27.8744 | -103.0276
27.8781 | -103.0327
27.8819 | -103.0391
27.8856 | -103.0453
27.8894 | -103.0533
27.8931 | -103.0613
27.8969 | -103.0693
27.9006 | -103.0765
27.9044 | -103.0837
27.9081 | -103.0907
27.9119 | -103.0978
27.9156 | -103.103
27.9194 | -103.1064
27.9231 | -103.1102
27.9269 | -103.1128
27.9306 | -103.1169
27.9344 | -103.1204
27.9381 | -103.1232
27.9419 | -103.1257
27.9456 | -103.1287
27.9494 | -103.1314
27.9531 | -103.1368
27.9569 | -103.1423
27.9606 | -103.1455
27.9644 | -103.1479
27.9681 | -103.1507
27.9719 | -103.1531
27.9756 | -103.1549
27.9794 | -103.1589
27.9831 | -103.1615
27.9869 | -103.1614
27.9906 | -103.1638
27.9944 | -103.1665
27.9981 | -103.169
28.0019 | -103.172