## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 254 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 204
Hood | 02-92B/P | 160
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1021.4 | hPa |  | PRC PRO 01
Humidity | 52.5 | %/Hr
Internal temperature of calibration bench | 23.34 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1025 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.6 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.09 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.06 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3100 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 22/06/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 28/06/2018
Technical Manager | Anthony HUE | Approval |  | 16/07/2017
CEO | Robert GLEMOT | Approval |  | 16/07/2017
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.2945 | 57.3769
0.012209 | -35.8478 | 38.3062
0.018313 | -34.8107 | 28.3666
0.024417 | -34.3624 | 22.0056
0.030521 | -34.2595 | 18.2317
0.036626 | -34.1188 | 15.1481
0.04273 | -34.0356 | 12.842
0.048834 | -33.9833 | 11.2615
0.054938 | -33.9142 | 10.0647
0.061043 | -33.8403 | 8.9443
0.067147 | -33.8299 | 8.0798
0.073251 | -33.8307 | 7.326
0.079355 | -33.8481 | 6.7656
0.08546 | -33.8657 | 6.3092
0.091564 | -33.8619 | 5.937
0.097668 | -33.8277 | 5.6806
0.10377 | -33.7803 | 5.3332
0.10988 | -33.7496 | 4.9719
0.11598 | -33.7406 | 4.6152
0.12209 | -33.7574 | 4.3396
0.12819 | -33.79 | 4.1098
0.13429 | -33.8217 | 3.8718
0.1404 | -33.8388 | 3.5998
0.1465 | -33.8545 | 3.343
0.15261 | -33.8618 | 3.105
0.15871 | -33.8587 | 2.9095
0.16482 | -33.8446 | 2.7503
0.17092 | -33.82 | 2.6854
0.17702 | -33.801 | 2.5859
0.18313 | -33.7836 | 2.5015
0.18923 | -33.771 | 2.3953
0.19534 | -33.7642 | 2.2931
0.20144 | -33.7633 | 2.1821
0.20754 | -33.7666 | 2.111
0.21365 | -33.7693 | 2.0383
0.21975 | -33.774 | 1.9751
0.22586 | -33.7823 | 1.9112
0.23196 | -33.7909 | 1.8388
0.23807 | -33.7995 | 1.7609
0.24417 | -33.8014 | 1.6669
0.25027 | -33.8045 | 1.5822
0.25638 | -33.8042 | 1.4895
0.26248 | -33.802 | 1.4021
0.26859 | -33.7994 | 1.3184
0.27469 | -33.7985 | 1.2315
0.2808 | -33.8042 | 1.1677
0.2869 | -33.8085 | 1.0996
0.293 | -33.8138 | 1.0427
0.29911 | -33.8199 | 1.0047
0.30521 | -33.8258 | 0.96867
0.31132 | -33.8298 | 0.95521
0.31742 | -33.8341 | 0.93325
0.32353 | -33.8376 | 0.91521
0.32963 | -33.8405 | 0.89998
0.33573 | -33.8424 | 0.88631
0.34184 | -33.845 | 0.86923
0.34794 | -33.8483 | 0.83308
0.35405 | -33.8516 | 0.80235
0.36015 | -33.8552 | 0.76792
0.36626 | -33.8575 | 0.73009
0.37236 | -33.86 | 0.67971
0.37846 | -33.8622 | 0.62225
0.38457 | -33.8611 | 0.56375
0.39067 | -33.8617 | 0.51373
0.39678 | -33.8599 | 0.45786
0.40288 | -33.8575 | 0.40541
0.40899 | -33.8537 | 0.35413
0.41509 | -33.8491 | 0.30267
0.42119 | -33.8441 | 0.25601
0.4273 | -33.8402 | 0.2077
0.4334 | -33.8367 | 0.15868
0.43951 | -33.8336 | 0.11037
0.44561 | -33.832 | 0.065059
0.45172 | -33.8335 | 0.027238
0.45782 | -33.8332 | -0.014503
0.46392 | -33.8344 | -0.052895
0.47003 | -33.836 | -0.087956
0.47613 | -33.8381 | -0.11989
0.48224 | -33.8403 | -0.15023
0.48834 | -33.8429 | -0.17134
0.49445 | -33.845 | -0.19572
0.50055 | -33.8477 | -0.21788
0.50665 | -33.8495 | -0.23568
0.51276 | -33.8508 | -0.25497
0.51886 | -33.8521 | -0.27172
0.52497 | -33.8513 | -0.27785
0.53107 | -33.8518 | -0.28752
0.53717 | -33.8522 | -0.29823
0.54328 | -33.8521 | -0.30306
0.54938 | -33.8521 | -0.31096
0.55549 | -33.8517 | -0.31988
0.56159 | -33.8516 | -0.33156
0.5677 | -33.8512 | -0.34234
0.5738 | -33.8507 | -0.35595
0.5799 | -33.8507 | -0.36845
0.58601 | -33.8511 | -0.3857
0.59211 | -33.8521 | -0.40934
0.59822 | -33.8527 | -0.42677
0.60432 | -33.8537 | -0.44901
0.61043 | -33.8547 | -0.47119
0.61653 | -33.8559 | -0.49347
0.62263 | -33.8569 | -0.51736
0.62874 | -33.8583 | -0.53817
0.63484 | -33.8595 | -0.55994
0.64095 | -33.8607 | -0.58154
0.64705 | -33.8619 | -0.60212
0.65316 | -33.863 | -0.62204
0.65926 | -33.8641 | -0.64475
0.66536 | -33.8636 | -0.66311
0.67147 | -33.864 | -0.67972
0.67757 | -33.8646 | -0.69785
0.68368 | -33.8645 | -0.71782
0.68978 | -33.8647 | -0.73593
0.69589 | -33.8644 | -0.75423
0.70199 | -33.8637 | -0.776
0.70809 | -33.8631 | -0.7962
0.7142 | -33.8627 | -0.81543
0.7203 | -33.862 | -0.83504
0.72641 | -33.8612 | -0.85582
0.73251 | -33.861 | -0.87666
0.73862 | -33.861 | -0.90103
0.74472 | -33.8613 | -0.92041
0.75082 | -33.8609 | -0.9414
0.75693 | -33.861 | -0.96494
0.76303 | -33.8609 | -0.98746
0.76914 | -33.8614 | -1.0147
0.77524 | -33.8615 | -1.0398
0.78135 | -33.8619 | -1.0633
0.78745 | -33.8624 | -1.0883
0.79355 | -33.8629 | -1.1131
0.79966 | -33.8638 | -1.139
0.80576 | -33.8648 | -1.1697
0.81187 | -33.8651 | -1.1993
0.81797 | -33.8656 | -1.2302
0.82408 | -33.8662 | -1.2575
0.83018 | -33.8667 | -1.2889
0.83628 | -33.8675 | -1.3197
0.84239 | -33.8673 | -1.3479
0.84849 | -33.8672 | -1.3761
0.8546 | -33.8676 | -1.404
0.8607 | -33.8676 | -1.4334
0.86681 | -33.8676 | -1.462
0.87291 | -33.8672 | -1.4878
0.87901 | -33.8666 | -1.5103
0.88512 | -33.8665 | -1.5334
0.89122 | -33.8663 | -1.5569
0.89733 | -33.866 | -1.5764
0.90343 | -33.8657 | -1.5977
0.90953 | -33.8659 | -1.6127
0.91564 | -33.8654 | -1.6287
0.92174 | -33.865 | -1.6431
0.92785 | -33.8645 | -1.6567
0.93395 | -33.8646 | -1.6685
0.94006 | -33.8646 | -1.6823
0.94616 | -33.8644 | -1.687
0.95226 | -33.8645 | -1.6986
0.95837 | -33.8647 | -1.7103
0.96447 | -33.8651 | -1.7206
0.97058 | -33.8651 | -1.7306
0.97668 | -33.8649 | -1.7386
0.98279 | -33.8649 | -1.7423
0.98889 | -33.8652 | -1.7472
0.99499 | -33.8653 | -1.7557
1.0011 | -33.8656 | -1.7607
1.0072 | -33.8655 | -1.7645
1.0133 | -33.8654 | -1.7685
1.0194 | -33.8655 | -1.7739
1.0255 | -33.8652 | -1.7808
1.0316 | -33.8646 | -1.785
1.0377 | -33.8641 | -1.7909
1.0438 | -33.8634 | -1.7954
1.0499 | -33.8619 | -1.8017
1.056 | -33.8607 | -1.8065
1.0621 | -33.8596 | -1.8122
1.0682 | -33.8591 | -1.8172
1.0743 | -33.858 | -1.8249
1.0805 | -33.8574 | -1.8317
1.0866 | -33.8558 | -1.8425
1.0927 | -33.8545 | -1.8499
1.0988 | -33.8537 | -1.8605
1.1049 | -33.8522 | -1.868
1.111 | -33.8507 | -1.8791
1.1171 | -33.8493 | -1.8884
1.1232 | -33.848 | -1.9027
1.1293 | -33.8468 | -1.9146
1.1354 | -33.8459 | -1.927
1.1415 | -33.8447 | -1.9411
1.1476 | -33.8433 | -1.955
1.1537 | -33.842 | -1.9683
1.1598 | -33.8418 | -1.9867
1.1659 | -33.8413 | -1.9988
1.172 | -33.8406 | -2.0122
1.1781 | -33.8399 | -2.0263
1.1842 | -33.8393 | -2.042
1.1903 | -33.8401 | -2.0573
1.1964 | -33.84 | -2.0719
1.2025 | -33.8399 | -2.0872
1.2086 | -33.8404 | -2.1031
1.2147 | -33.8406 | -2.1182
1.2209 | -33.8414 | -2.1318
1.227 | -33.8433 | -2.1471
1.2331 | -33.8443 | -2.1615
1.2392 | -33.8452 | -2.1775
1.2453 | -33.8465 | -2.1898
1.2514 | -33.8478 | -2.2021
1.2575 | -33.8494 | -2.2164
1.2636 | -33.8518 | -2.2281
1.2697 | -33.8535 | -2.2401
1.2758 | -33.8551 | -2.2544
1.2819 | -33.8569 | -2.2682
1.288 | -33.8586 | -2.2822
1.2941 | -33.8607 | -2.2932
1.3002 | -33.8627 | -2.3018
1.3063 | -33.8644 | -2.312
1.3124 | -33.8663 | -2.3223
1.3185 | -33.8683 | -2.3332
1.3246 | -33.8701 | -2.3443
1.3307 | -33.8723 | -2.354
1.3368 | -33.8741 | -2.3618
1.3429 | -33.8758 | -2.3709
1.349 | -33.8773 | -2.3796
1.3551 | -33.879 | -2.3908
1.3613 | -33.8806 | -2.3989
1.3674 | -33.8819 | -2.4076
1.3735 | -33.8832 | -2.4159
1.3796 | -33.8845 | -2.4235
1.3857 | -33.8855 | -2.4308
1.3918 | -33.8866 | -2.4407
1.3979 | -33.8879 | -2.4494
1.404 | -33.8886 | -2.4596
1.4101 | -33.8896 | -2.4687
1.4162 | -33.8904 | -2.4775
1.4223 | -33.8912 | -2.488
1.4284 | -33.892 | -2.4965
1.4345 | -33.8927 | -2.507
1.4406 | -33.8932 | -2.5182
1.4467 | -33.8939 | -2.5277
1.4528 | -33.8945 | -2.5371
1.4589 | -33.8954 | -2.5477
1.465 | -33.8956 | -2.5556
1.4711 | -33.8966 | -2.5647
1.4772 | -33.8971 | -2.577
1.4833 | -33.8977 | -2.5893
1.4894 | -33.8985 | -2.6009
1.4955 | -33.8992 | -2.6125
1.5016 | -33.8999 | -2.6247
1.5078 | -33.9003 | -2.6394
1.5139 | -33.901 | -2.6521
1.52 | -33.9015 | -2.6655
1.5261 | -33.902 | -2.6773
1.5322 | -33.9022 | -2.69
1.5383 | -33.9025 | -2.7023
1.5444 | -33.9032 | -2.7158
1.5505 | -33.9036 | -2.7292
1.5566 | -33.9041 | -2.7422
1.5627 | -33.9047 | -2.7556
1.5688 | -33.9055 | -2.7688
1.5749 | -33.9058 | -2.7813
1.581 | -33.9062 | -2.7981
1.5871 | -33.9067 | -2.8139
1.5932 | -33.9072 | -2.827
1.5993 | -33.9075 | -2.8439
1.6054 | -33.9079 | -2.859
1.6115 | -33.9082 | -2.8761
1.6176 | -33.9086 | -2.8927
1.6237 | -33.9089 | -2.9072
1.6298 | -33.9092 | -2.9228
1.6359 | -33.9096 | -2.9375
1.642 | -33.91 | -2.9523
1.6482 | -33.9104 | -2.9695
1.6543 | -33.9109 | -2.9864
1.6604 | -33.9113 | -3.0013
1.6665 | -33.9114 | -3.0164
1.6726 | -33.9117 | -3.0329
1.6787 | -33.9123 | -3.0501
1.6848 | -33.9121 | -3.0667
1.6909 | -33.9125 | -3.0832
1.697 | -33.9127 | -3.0996
1.7031 | -33.9128 | -3.1155
1.7092 | -33.9128 | -3.1307
1.7153 | -33.913 | -3.1471
1.7214 | -33.9126 | -3.1627
1.7275 | -33.9124 | -3.1788
1.7336 | -33.9122 | -3.1952
1.7397 | -33.9122 | -3.2107
1.7458 | -33.9123 | -3.2276
1.7519 | -33.9121 | -3.244
1.758 | -33.912 | -3.2591
1.7641 | -33.9121 | -3.2738
1.7702 | -33.9118 | -3.2898
1.7763 | -33.9119 | -3.3037
1.7824 | -33.9117 | -3.3201
1.7885 | -33.9115 | -3.3343
1.7947 | -33.9116 | -3.3494
1.8008 | -33.9115 | -3.3625
1.8069 | -33.9114 | -3.3764
1.813 | -33.9114 | -3.3908
1.8191 | -33.9111 | -3.4052
1.8252 | -33.9109 | -3.4193
1.8313 | -33.9109 | -3.4339
1.8374 | -33.9109 | -3.4476
1.8435 | -33.911 | -3.4599
1.8496 | -33.9107 | -3.4746
1.8557 | -33.9106 | -3.4895
1.8618 | -33.9106 | -3.5015
1.8679 | -33.9106 | -3.5147
1.874 | -33.9106 | -3.5281
1.8801 | -33.9104 | -3.5412
1.8862 | -33.9106 | -3.5542
1.8923 | -33.9105 | -3.5665
1.8984 | -33.9102 | -3.5773
1.9045 | -33.9101 | -3.5886
1.9106 | -33.9098 | -3.6006
1.9167 | -33.9095 | -3.612
1.9228 | -33.9095 | -3.623
1.9289 | -33.9092 | -3.6321
1.9351 | -33.9093 | -3.6413
1.9412 | -33.9093 | -3.6518
1.9473 | -33.9094 | -3.6639
1.9534 | -33.9093 | -3.6752
1.9595 | -33.909 | -3.6869
1.9656 | -33.9086 | -3.6977
1.9717 | -33.9083 | -3.7084
1.9778 | -33.9084 | -3.7201
1.9839 | -33.9084 | -3.7316
1.99 | -33.9084 | -3.7417
1.9961 | -33.9082 | -3.7523
2.0022 | -33.9081 | -3.7624
2.0083 | -33.9078 | -3.7725
2.0144 | -33.9078 | -3.7825
2.0205 | -33.9078 | -3.7945
2.0266 | -33.9079 | -3.8062
2.0327 | -33.9079 | -3.8176
2.0388 | -33.9082 | -3.8316
2.0449 | -33.9082 | -3.8421
2.051 | -33.9083 | -3.8523
2.0571 | -33.9082 | -3.8638
2.0632 | -33.9081 | -3.8745
2.0693 | -33.9081 | -3.8858
2.0754 | -33.9083 | -3.8978
2.0816 | -33.9084 | -3.9099
2.0877 | -33.9084 | -3.9211
2.0938 | -33.9086 | -3.9337
2.0999 | -33.9086 | -3.9454
2.106 | -33.9088 | -3.9575
2.1121 | -33.9088 | -3.97
2.1182 | -33.909 | -3.9802
2.1243 | -33.9092 | -3.991
2.1304 | -33.9093 | -4.0014
2.1365 | -33.9096 | -4.0131
2.1426 | -33.9098 | -4.0266
2.1487 | -33.9102 | -4.0385
2.1548 | -33.9107 | -4.0517
2.1609 | -33.9112 | -4.064
2.167 | -33.9117 | -4.0767
2.1731 | -33.912 | -4.0888
2.1792 | -33.9127 | -4.1002
2.1853 | -33.9132 | -4.1132
2.1914 | -33.9137 | -4.1256
2.1975 | -33.9143 | -4.138
2.2036 | -33.9148 | -4.1513
2.2097 | -33.9152 | -4.1635
2.2158 | -33.9161 | -4.174
2.222 | -33.9167 | -4.1843
2.2281 | -33.917 | -4.1953
2.2342 | -33.9176 | -4.2087
2.2403 | -33.918 | -4.219
2.2464 | -33.9188 | -4.2303
2.2525 | -33.9193 | -4.2412
2.2586 | -33.9202 | -4.2513
2.2647 | -33.921 | -4.2608
2.2708 | -33.9216 | -4.2727
2.2769 | -33.9225 | -4.2844
2.283 | -33.9229 | -4.2939
2.2891 | -33.9235 | -4.3044
2.2952 | -33.9241 | -4.3144
2.3013 | -33.9248 | -4.3248
2.3074 | -33.9255 | -4.3346
2.3135 | -33.9259 | -4.3461
2.3196 | -33.9266 | -4.3549
2.3257 | -33.9269 | -4.3656
2.3318 | -33.9273 | -4.3748
2.3379 | -33.9278 | -4.3864
2.344 | -33.9281 | -4.3975
2.3501 | -33.9286 | -4.4084
2.3562 | -33.9289 | -4.4169
2.3623 | -33.9292 | -4.4257
2.3685 | -33.9294 | -4.435
2.3746 | -33.9297 | -4.4456
2.3807 | -33.9302 | -4.4548
2.3868 | -33.9304 | -4.4639
2.3929 | -33.9308 | -4.4735
2.399 | -33.931 | -4.4816
2.4051 | -33.9315 | -4.4913
2.4112 | -33.9319 | -4.5005
2.4173 | -33.9322 | -4.5096
2.4234 | -33.9322 | -4.5182
2.4295 | -33.9325 | -4.5269
2.4356 | -33.9326 | -4.5364
2.4417 | -33.9326 | -4.5466
2.4478 | -33.9327 | -4.5563
2.4539 | -33.933 | -4.5644
2.46 | -33.9327 | -4.5742
2.4661 | -33.9327 | -4.5836
2.4722 | -33.933 | -4.5925
2.4783 | -33.9331 | -4.601
2.4844 | -33.9331 | -4.6097
2.4905 | -33.933 | -4.6191
2.4966 | -33.9327 | -4.6281
2.5027 | -33.9327 | -4.6364
2.5089 | -33.9326 | -4.645
2.515 | -33.9324 | -4.6544
2.5211 | -33.9323 | -4.6629
2.5272 | -33.932 | -4.6721
2.5333 | -33.9318 | -4.6818
2.5394 | -33.9318 | -4.691
2.5455 | -33.9316 | -4.7005
2.5516 | -33.9315 | -4.7094
2.5577 | -33.9315 | -4.7195
2.5638 | -33.9312 | -4.73
2.5699 | -33.931 | -4.7404
2.576 | -33.9308 | -4.7509
2.5821 | -33.9306 | -4.761
2.5882 | -33.9304 | -4.7707
2.5943 | -33.9303 | -4.7807
2.6004 | -33.9301 | -4.7909
2.6065 | -33.9299 | -4.7995
2.6126 | -33.9298 | -4.8088
2.6187 | -33.9297 | -4.8177
2.6248 | -33.9296 | -4.8269
2.6309 | -33.9294 | -4.8384
2.637 | -33.9294 | -4.8497
2.6431 | -33.9294 | -4.8594
2.6492 | -33.9293 | -4.8696
2.6554 | -33.9293 | -4.88
2.6615 | -33.9291 | -4.89
2.6676 | -33.929 | -4.9006
2.6737 | -33.9291 | -4.9116
2.6798 | -33.9292 | -4.9216
2.6859 | -33.9292 | -4.9324
2.692 | -33.9291 | -4.9423
2.6981 | -33.929 | -4.9536
2.7042 | -33.9291 | -4.9659
2.7103 | -33.9291 | -4.9762
2.7164 | -33.9293 | -4.9871
2.7225 | -33.9294 | -4.9986
2.7286 | -33.9295 | -5.01
2.7347 | -33.9295 | -5.0202
2.7408 | -33.9298 | -5.0297
2.7469 | -33.9298 | -5.0412
2.753 | -33.9301 | -5.0521
2.7591 | -33.9302 | -5.0642
2.7652 | -33.9302 | -5.076
2.7713 | -33.9302 | -5.086
2.7774 | -33.9307 | -5.0977
2.7835 | -33.9308 | -5.1086
2.7896 | -33.9309 | -5.1181
2.7958 | -33.9313 | -5.1286
2.8019 | -33.9319 | -5.1386
2.808 | -33.9322 | -5.1508
2.8141 | -33.9326 | -5.164
2.8202 | -33.933 | -5.1764
2.8263 | -33.9334 | -5.1876
2.8324 | -33.9337 | -5.2002
2.8385 | -33.9341 | -5.2126
2.8446 | -33.9347 | -5.2245
2.8507 | -33.935 | -5.2369
2.8568 | -33.9355 | -5.2483
2.8629 | -33.9359 | -5.2598
2.869 | -33.9364 | -5.2714
2.8751 | -33.9368 | -5.2837
2.8812 | -33.9375 | -5.2972
2.8873 | -33.9378 | -5.3088
2.8934 | -33.9382 | -5.3209
2.8995 | -33.9386 | -5.3331
2.9056 | -33.9389 | -5.3441
2.9117 | -33.9394 | -5.3559
2.9178 | -33.9401 | -5.3689
2.9239 | -33.9405 | -5.3825
2.93 | -33.941 | -5.3962
2.9361 | -33.9415 | -5.4079
2.9423 | -33.9421 | -5.4201
2.9484 | -33.9427 | -5.4339
2.9545 | -33.9434 | -5.4474
2.9606 | -33.9439 | -5.4611
2.9667 | -33.9444 | -5.4732
2.9728 | -33.9448 | -5.4857
2.9789 | -33.9453 | -5.4987
2.985 | -33.9457 | -5.5127
2.9911 | -33.9462 | -5.5255
2.9972 | -33.9468 | -5.5385
3.0033 | -33.9473 | -5.5518
3.0094 | -33.9479 | -5.5651
3.0155 | -33.9485 | -5.578
3.0216 | -33.949 | -5.5917
3.0277 | -33.9495 | -5.6056
3.0338 | -33.9501 | -5.6195
3.0399 | -33.9507 | -5.6334
3.046 | -33.951 | -5.6467
3.0521 | -33.9516 | -5.6594
3.0582 | -33.9519 | -5.6735
3.0643 | -33.9524 | -5.6868
3.0704 | -33.9527 | -5.7005
3.0765 | -33.9531 | -5.7129
3.0827 | -33.9534 | -5.7263
3.0888 | -33.9538 | -5.7394
3.0949 | -33.9541 | -5.7527
3.101 | -33.9546 | -5.767
3.1071 | -33.9549 | -5.7814
3.1132 | -33.9553 | -5.7957
3.1193 | -33.9558 | -5.8091
3.1254 | -33.9562 | -5.8215
3.1315 | -33.9566 | -5.8361
3.1376 | -33.9568 | -5.8492
3.1437 | -33.9572 | -5.863
3.1498 | -33.9576 | -5.8775
3.1559 | -33.9579 | -5.8909
3.162 | -33.9579 | -5.9052
3.1681 | -33.958 | -5.9197
3.1742 | -33.9584 | -5.9336
3.1803 | -33.9588 | -5.9471
3.1864 | -33.9592 | -5.9596
3.1925 | -33.9595 | -5.9735
3.1986 | -33.9596 | -5.9865
3.2047 | -33.9597 | -6.0007
3.2108 | -33.9599 | -6.0127
3.2169 | -33.9604 | -6.0267
3.223 | -33.9603 | -6.0403
3.2292 | -33.9603 | -6.0545
3.2353 | -33.9602 | -6.0642
3.2414 | -33.9604 | -6.0774
3.2475 | -33.9604 | -6.0887
3.2536 | -33.9607 | -6.1027
3.2597 | -33.9609 | -6.1141
3.2658 | -33.9611 | -6.1251
3.2719 | -33.9612 | -6.1363
3.278 | -33.9613 | -6.1487
3.2841 | -33.9614 | -6.1608
3.2902 | -33.9615 | -6.1736
3.2963 | -33.9617 | -6.1855
3.3024 | -33.9616 | -6.1958
3.3085 | -33.9618 | -6.2086
3.3146 | -33.9617 | -6.2197
3.3207 | -33.9619 | -6.2321
3.3268 | -33.9621 | -6.2445
3.3329 | -33.9622 | -6.2545
3.339 | -33.9623 | -6.2643
3.3451 | -33.9623 | -6.2759
3.3512 | -33.9624 | -6.2856
3.3573 | -33.9626 | -6.2971
3.3634 | -33.9627 | -6.3088
3.3696 | -33.9627 | -6.3209
3.3757 | -33.9627 | -6.3309
3.3818 | -33.963 | -6.3428
3.3879 | -33.9629 | -6.3533
3.394 | -33.9629 | -6.364
3.4001 | -33.9629 | -6.3751
3.4062 | -33.963 | -6.3854
3.4123 | -33.963 | -6.3954
3.4184 | -33.9631 | -6.4066
3.4245 | -33.9632 | -6.4173
3.4306 | -33.9631 | -6.4272
3.4367 | -33.9633 | -6.4372
3.4428 | -33.9636 | -6.4465
3.4489 | -33.9635 | -6.4578
3.455 | -33.9634 | -6.4684
3.4611 | -33.9635 | -6.4797
3.4672 | -33.9635 | -6.4912
3.4733 | -33.9636 | -6.5015
3.4794 | -33.9636 | -6.5122
3.4855 | -33.9638 | -6.5219
3.4916 | -33.9639 | -6.5318
3.4977 | -33.9639 | -6.5429
3.5038 | -33.9641 | -6.553
3.5099 | -33.9641 | -6.5628
3.5161 | -33.9643 | -6.5725
3.5222 | -33.9644 | -6.5827
3.5283 | -33.9643 | -6.5929
3.5344 | -33.9644 | -6.6038
3.5405 | -33.9645 | -6.6139
3.5466 | -33.9647 | -6.6239
3.5527 | -33.9648 | -6.6347
3.5588 | -33.9649 | -6.6446
3.5649 | -33.9649 | -6.6549
3.571 | -33.965 | -6.6653
3.5771 | -33.9651 | -6.6762
3.5832 | -33.9653 | -6.686
3.5893 | -33.9655 | -6.696
3.5954 | -33.9656 | -6.7066
3.6015 | -33.9658 | -6.7172
3.6076 | -33.966 | -6.7269
3.6137 | -33.9661 | -6.737
3.6198 | -33.9665 | -6.7462
3.6259 | -33.9665 | -6.7571
3.632 | -33.9668 | -6.7679
3.6381 | -33.967 | -6.7779
3.6442 | -33.9672 | -6.7885
3.6503 | -33.9675 | -6.7977
3.6565 | -33.9678 | -6.8096
3.6626 | -33.9679 | -6.8205
3.6687 | -33.9682 | -6.8313
3.6748 | -33.9685 | -6.8413
3.6809 | -33.9689 | -6.852
3.687 | -33.9691 | -6.8633
3.6931 | -33.9695 | -6.8735
3.6992 | -33.9697 | -6.8843
3.7053 | -33.9699 | -6.8946
3.7114 | -33.9701 | -6.906
3.7175 | -33.9705 | -6.916
3.7236 | -33.9707 | -6.9251
3.7297 | -33.9708 | -6.936
3.7358 | -33.9711 | -6.9478
3.7419 | -33.9713 | -6.9596
3.748 | -33.9715 | -6.9705
3.7541 | -33.9716 | -6.9817
3.7602 | -33.972 | -6.9927
3.7663 | -33.9722 | -7.003
3.7724 | -33.9726 | -7.0142
3.7785 | -33.9729 | -7.0248
3.7846 | -33.9731 | -7.0356
3.7907 | -33.9736 | -7.0468
3.7969 | -33.9738 | -7.0585
3.803 | -33.9741 | -7.0687
3.8091 | -33.9745 | -7.0792
3.8152 | -33.9748 | -7.0906
3.8213 | -33.975 | -7.1024
3.8274 | -33.9754 | -7.1131
3.8335 | -33.9758 | -7.1249
3.8396 | -33.9762 | -7.1366
3.8457 | -33.9766 | -7.147
3.8518 | -33.9771 | -7.1581
3.8579 | -33.9775 | -7.1688
3.864 | -33.9779 | -7.1797
3.8701 | -33.9783 | -7.1909
3.8762 | -33.9786 | -7.2022
3.8823 | -33.9789 | -7.213
3.8884 | -33.9793 | -7.2236
3.8945 | -33.9798 | -7.2351
3.9006 | -33.9802 | -7.2468
3.9067 | -33.9806 | -7.2588
3.9128 | -33.981 | -7.2703
3.9189 | -33.9814 | -7.2811
3.925 | -33.9817 | -7.2922
3.9311 | -33.9821 | -7.3024
3.9372 | -33.9825 | -7.3142
3.9434 | -33.9829 | -7.3248
3.9495 | -33.9833 | -7.3366
3.9556 | -33.9837 | -7.348
3.9617 | -33.9841 | -7.3591
3.9678 | -33.9845 | -7.37
3.9739 | -33.9849 | -7.3814
3.98 | -33.9853 | -7.3935
3.9861 | -33.9856 | -7.4041
3.9922 | -33.9858 | -7.416
3.9983 | -33.9862 | -7.4263
4.0044 | -33.9866 | -7.4362
4.0105 | -33.9869 | -7.4473
4.0166 | -33.9875 | -7.4585
4.0227 | -33.9879 | -7.4691
4.0288 | -33.9881 | -7.4806
4.0349 | -33.9885 | -7.4907
4.041 | -33.9888 | -7.5018
4.0471 | -33.9891 | -7.5123
4.0532 | -33.9895 | -7.5236
4.0593 | -33.9898 | -7.5354
4.0654 | -33.9902 | -7.5464
4.0715 | -33.9905 | -7.558
4.0776 | -33.9909 | -7.5685
4.0838 | -33.9913 | -7.5801
4.0899 | -33.9916 | -7.5916
4.096 | -33.9919 | -7.6022
4.1021 | -33.9923 | -7.6132
4.1082 | -33.9927 | -7.6237
4.1143 | -33.993 | -7.6338
4.1204 | -33.9935 | -7.6448
4.1265 | -33.9939 | -7.6554
4.1326 | -33.9942 | -7.6659
4.1387 | -33.9947 | -7.6774
4.1448 | -33.995 | -7.6885
4.1509 | -33.9953 | -7.6994
4.157 | -33.9956 | -7.7108
4.1631 | -33.9959 | -7.7223
4.1692 | -33.9962 | -7.7332
4.1753 | -33.9966 | -7.7442
4.1814 | -33.9969 | -7.7542
4.1875 | -33.9972 | -7.7658
4.1936 | -33.9975 | -7.7773
4.1997 | -33.9979 | -7.7877
4.2058 | -33.9981 | -7.7994
4.2119 | -33.9984 | -7.8094
4.218 | -33.9987 | -7.8197
4.2241 | -33.9991 | -7.8295
4.2303 | -33.9995 | -7.8409
4.2364 | -33.9998 | -7.8525
4.2425 | -34.0001 | -7.8633
4.2486 | -34.0004 | -7.8746
4.2547 | -34.0005 | -7.884
4.2608 | -34.0008 | -7.8941
4.2669 | -34.0011 | -7.9047
4.273 | -34.0014 | -7.9157
4.2791 | -34.0017 | -7.9266
4.2852 | -34.0019 | -7.9371
4.2913 | -34.002 | -7.9465
4.2974 | -34.0023 | -7.9576
4.3035 | -34.0025 | -7.9687
4.3096 | -34.0027 | -7.9793
4.3157 | -34.003 | -7.9895
4.3218 | -34.0032 | -7.9996
4.3279 | -34.0034 | -8.0102
4.334 | -34.0037 | -8.0202
4.3401 | -34.0039 | -8.0304
4.3462 | -34.0042 | -8.0414
4.3523 | -34.0046 | -8.0512
4.3584 | -34.0047 | -8.0599
4.3645 | -34.0049 | -8.0697
4.3707 | -34.0051 | -8.0792
4.3768 | -34.0053 | -8.0898
4.3829 | -34.0055 | -8.1003
4.389 | -34.0058 | -8.1101
4.3951 | -34.006 | -8.1198
4.4012 | -34.0062 | -8.1298
4.4073 | -34.0065 | -8.1401
4.4134 | -34.0067 | -8.1505
4.4195 | -34.007 | -8.1603
4.4256 | -34.0073 | -8.1701
4.4317 | -34.0075 | -8.1797
4.4378 | -34.0077 | -8.1895
4.4439 | -34.0079 | -8.1991
4.45 | -34.008 | -8.2092
4.4561 | -34.0081 | -8.2193
4.4622 | -34.0083 | -8.229
4.4683 | -34.0085 | -8.2384
4.4744 | -34.0087 | -8.2487
4.4805 | -34.009 | -8.258
4.4866 | -34.0092 | -8.2674
4.4927 | -34.0094 | -8.2778
4.4988 | -34.0096 | -8.2874
4.5049 | -34.0098 | -8.2967
4.511 | -34.0101 | -8.3059
4.5172 | -34.0104 | -8.3159
4.5233 | -34.0106 | -8.3267
4.5294 | -34.0108 | -8.3362
4.5355 | -34.011 | -8.3458
4.5416 | -34.0113 | -8.355
4.5477 | -34.0115 | -8.3645
4.5538 | -34.0117 | -8.3734
4.5599 | -34.012 | -8.3831
4.566 | -34.0122 | -8.393
4.5721 | -34.0125 | -8.4023
4.5782 | -34.0128 | -8.4123
4.5843 | -34.013 | -8.4221
4.5904 | -34.0133 | -8.4315
4.5965 | -34.0136 | -8.4405
4.6026 | -34.0139 | -8.4512
4.6087 | -34.0141 | -8.4604
4.6148 | -34.0143 | -8.4712
4.6209 | -34.0147 | -8.4799
4.627 | -34.0151 | -8.4901
4.6331 | -34.0151 | -8.4991
4.6392 | -34.0154 | -8.5085
4.6453 | -34.0156 | -8.5181
4.6514 | -34.0159 | -8.5278
4.6576 | -34.016 | -8.5384
4.6637 | -34.0162 | -8.5462
4.6698 | -34.0165 | -8.555
4.6759 | -34.0166 | -8.5649
4.682 | -34.0168 | -8.5742
4.6881 | -34.017 | -8.5832
4.6942 | -34.0172 | -8.5926
4.7003 | -34.0175 | -8.6022
4.7064 | -34.0177 | -8.6117
4.7125 | -34.0181 | -8.6213
4.7186 | -34.0182 | -8.6299
4.7247 | -34.0186 | -8.6397
4.7308 | -34.0189 | -8.6491
4.7369 | -34.0191 | -8.6589
4.743 | -34.0193 | -8.6693
4.7491 | -34.0195 | -8.6791
4.7552 | -34.0196 | -8.6885
4.7613 | -34.0198 | -8.6982
4.7674 | -34.0201 | -8.707
4.7735 | -34.0202 | -8.7172
4.7796 | -34.0205 | -8.7272
4.7857 | -34.0206 | -8.737
4.7918 | -34.0209 | -8.7458
4.7979 | -34.0212 | -8.755
4.8041 | -34.0213 | -8.7645
4.8102 | -34.0217 | -8.7738
4.8163 | -34.0219 | -8.7834
4.8224 | -34.022 | -8.7922
4.8285 | -34.0224 | -8.8014
4.8346 | -34.0226 | -8.8111
4.8407 | -34.0228 | -8.8206
4.8468 | -34.0231 | -8.8305
4.8529 | -34.0233 | -8.8404
4.859 | -34.0235 | -8.8501
4.8651 | -34.0238 | -8.8596
4.8712 | -34.024 | -8.8691
4.8773 | -34.0242 | -8.8789
4.8834 | -34.0245 | -8.8888
4.8895 | -34.0248 | -8.899
4.8956 | -34.025 | -8.9091
4.9017 | -34.0251 | -8.9188
4.9078 | -34.0254 | -8.9288
4.9139 | -34.0256 | -8.9376
4.92 | -34.0259 | -8.9468
4.9261 | -34.0262 | -8.9558
4.9322 | -34.0266 | -8.9651
4.9383 | -34.0269 | -8.9743
4.9445 | -34.0272 | -8.9846
4.9506 | -34.0275 | -8.9936
4.9567 | -34.0277 | -9.0031
4.9628 | -34.028 | -9.0126
4.9689 | -34.0282 | -9.0226
4.975 | -34.0284 | -9.0321
4.9811 | -34.0287 | -9.0421
4.9872 | -34.029 | -9.052
4.9933 | -34.0294 | -9.0619
4.9994 | -34.0297 | -9.0719
5.0055 | -34.0301 | -9.0817
5.0116 | -34.0304 | -9.0921
5.0177 | -34.0308 | -9.1024
5.0238 | -34.0311 | -9.1131
5.0299 | -34.0316 | -9.1243
5.036 | -34.0319 | -9.134
5.0421 | -34.0321 | -9.1438
5.0482 | -34.0324 | -9.1543
5.0543 | -34.0327 | -9.1638
5.0604 | -34.0331 | -9.1742
5.0665 | -34.0335 | -9.1841
5.0726 | -34.0338 | -9.1947
5.0787 | -34.0341 | -9.2051
5.0848 | -34.0344 | -9.2154
5.091 | -34.0347 | -9.2254
5.0971 | -34.0351 | -9.2357
5.1032 | -34.0355 | -9.2463
5.1093 | -34.0358 | -9.2564
5.1154 | -34.036 | -9.2666
5.1215 | -34.0363 | -9.2768
5.1276 | -34.0366 | -9.2868
5.1337 | -34.0371 | -9.2967
5.1398 | -34.0374 | -9.3067
5.1459 | -34.0378 | -9.3173
5.152 | -34.0381 | -9.3286
5.1581 | -34.0383 | -9.3394
5.1642 | -34.0386 | -9.3496
5.1703 | -34.039 | -9.3598
5.1764 | -34.0393 | -9.37
5.1825 | -34.0396 | -9.3805
5.1886 | -34.0399 | -9.3905
5.1947 | -34.0402 | -9.4015
5.2008 | -34.0407 | -9.4116
5.2069 | -34.0411 | -9.4229
5.213 | -34.0414 | -9.4327
5.2191 | -34.0418 | -9.4433
5.2252 | -34.0421 | -9.4538
5.2314 | -34.0423 | -9.4647
5.2375 | -34.0428 | -9.4748
5.2436 | -34.043 | -9.4852
5.2497 | -34.0435 | -9.4953
5.2558 | -34.0438 | -9.5065
5.2619 | -34.0441 | -9.5175
5.268 | -34.0445 | -9.5282
5.2741 | -34.0449 | -9.5396
5.2802 | -34.0453 | -9.5506
5.2863 | -34.0457 | -9.561
5.2924 | -34.0459 | -9.5724
5.2985 | -34.0464 | -9.584
5.3046 | -34.0469 | -9.5954
5.3107 | -34.0473 | -9.6065
5.3168 | -34.0477 | -9.6171
5.3229 | -34.0481 | -9.6272
5.329 | -34.0485 | -9.6373
5.3351 | -34.0489 | -9.6478
5.3412 | -34.0493 | -9.6581
5.3473 | -34.0497 | -9.6692
5.3534 | -34.0502 | -9.6803
5.3595 | -34.0505 | -9.6913
5.3656 | -34.0509 | -9.7025
5.3717 | -34.0512 | -9.7132
5.3779 | -34.0516 | -9.7244
5.384 | -34.0519 | -9.7353
5.3901 | -34.0523 | -9.7463
5.3962 | -34.0526 | -9.7573
5.4023 | -34.053 | -9.7683
5.4084 | -34.0534 | -9.7789
5.4145 | -34.0538 | -9.7902
5.4206 | -34.054 | -9.8017
5.4267 | -34.0545 | -9.8133
5.4328 | -34.0549 | -9.8236
5.4389 | -34.0553 | -9.8341
5.445 | -34.0555 | -9.8456
5.4511 | -34.056 | -9.8562
5.4572 | -34.0564 | -9.8677
5.4633 | -34.0568 | -9.8783
5.4694 | -34.057 | -9.889
5.4755 | -34.0575 | -9.9001
5.4816 | -34.0578 | -9.9109
5.4877 | -34.0582 | -9.9219
5.4938 | -34.0585 | -9.9325
5.4999 | -34.0589 | -9.9435
5.506 | -34.0593 | -9.9544
5.5121 | -34.0597 | -9.9657
5.5183 | -34.06 | -9.9761
5.5244 | -34.0603 | -9.9871
5.5305 | -34.0607 | -9.9982
5.5366 | -34.061 | -10.0089
5.5427 | -34.0614 | -10.0204
5.5488 | -34.0617 | -10.0317
5.5549 | -34.062 | -10.0426
5.561 | -34.0624 | -10.0541
5.5671 | -34.0628 | -10.0647
5.5732 | -34.0631 | -10.0759
5.5793 | -34.0635 | -10.0876
5.5854 | -34.064 | -10.0988
5.5915 | -34.0642 | -10.11
5.5976 | -34.0645 | -10.1212
5.6037 | -34.0649 | -10.1321
5.6098 | -34.0652 | -10.1429
5.6159 | -34.0656 | -10.154
5.622 | -34.066 | -10.1647
5.6281 | -34.0663 | -10.1758
5.6342 | -34.0667 | -10.1869
5.6403 | -34.067 | -10.1974
5.6464 | -34.0674 | -10.2084
5.6525 | -34.0677 | -10.2191
5.6586 | -34.0681 | -10.2293
5.6648 | -34.0684 | -10.2408
5.6709 | -34.0688 | -10.2521
5.677 | -34.0692 | -10.2639
5.6831 | -34.0695 | -10.2743
5.6892 | -34.0699 | -10.2853
5.6953 | -34.0702 | -10.2962
5.7014 | -34.0706 | -10.3077
5.7075 | -34.0709 | -10.3189
5.7136 | -34.0712 | -10.3297
5.7197 | -34.0716 | -10.3405
5.7258 | -34.0719 | -10.3513
5.7319 | -34.0722 | -10.3622
5.738 | -34.0726 | -10.3731
5.7441 | -34.073 | -10.384
5.7502 | -34.0734 | -10.3946
5.7563 | -34.0737 | -10.4053
5.7624 | -34.074 | -10.417
5.7685 | -34.0743 | -10.4277
5.7746 | -34.0746 | -10.4385
5.7807 | -34.0749 | -10.449
5.7868 | -34.0752 | -10.4599
5.7929 | -34.0755 | -10.4707
5.799 | -34.0759 | -10.4815
5.8052 | -34.0762 | -10.4925
5.8113 | -34.0765 | -10.5034
5.8174 | -34.0768 | -10.5144
5.8235 | -34.0771 | -10.5251
5.8296 | -34.0774 | -10.5366
5.8357 | -34.0777 | -10.5474
5.8418 | -34.078 | -10.5585
5.8479 | -34.0782 | -10.5697
5.854 | -34.0786 | -10.5808
5.8601 | -34.0789 | -10.5915
5.8662 | -34.0792 | -10.6027
5.8723 | -34.0795 | -10.6129
5.8784 | -34.0798 | -10.6237
5.8845 | -34.0802 | -10.6347
5.8906 | -34.0804 | -10.6455
5.8967 | -34.0808 | -10.656
5.9028 | -34.0812 | -10.6669
5.9089 | -34.0814 | -10.6771
5.915 | -34.0818 | -10.688
5.9211 | -34.0821 | -10.6988
5.9272 | -34.0824 | -10.7095
5.9333 | -34.0827 | -10.7197
5.9394 | -34.0831 | -10.7303
5.9455 | -34.0834 | -10.7406
5.9517 | -34.0838 | -10.7514
5.9578 | -34.0842 | -10.7625
5.9639 | -34.0845 | -10.7729
5.97 | -34.0849 | -10.7838
5.9761 | -34.0852 | -10.7943
5.9822 | -34.0855 | -10.8051
5.9883 | -34.0859 | -10.8156
5.9944 | -34.0862 | -10.8264
6.0005 | -34.0867 | -10.8374
6.0066 | -34.0869 | -10.8477
6.0127 | -34.0873 | -10.8578
6.0188 | -34.0876 | -10.8688
6.0249 | -34.0879 | -10.8795
6.031 | -34.0882 | -10.8899
6.0371 | -34.0886 | -10.9009
6.0432 | -34.089 | -10.9116
6.0493 | -34.0894 | -10.922
6.0554 | -34.0898 | -10.9326
6.0615 | -34.0901 | -10.9438
6.0676 | -34.0904 | -10.9551
6.0737 | -34.0908 | -10.966
6.0798 | -34.091 | -10.9765
6.0859 | -34.0914 | -10.9871
6.0921 | -34.0918 | -10.9975
6.0982 | -34.0922 | -11.0084
6.1043 | -34.0926 | -11.0189
6.1104 | -34.093 | -11.0301
6.1165 | -34.0933 | -11.04
6.1226 | -34.0937 | -11.05
6.1287 | -34.0941 | -11.0608
6.1348 | -34.0945 | -11.0715
6.1409 | -34.0948 | -11.0819
6.147 | -34.0952 | -11.0924
6.1531 | -34.0955 | -11.103
6.1592 | -34.0959 | -11.1136
6.1653 | -34.0962 | -11.124
6.1714 | -34.0966 | -11.1346
6.1775 | -34.097 | -11.1455
6.1836 | -34.0973 | -11.1559
6.1897 | -34.0976 | -11.1665
6.1958 | -34.0979 | -11.1773
6.2019 | -34.0983 | -11.1882
6.208 | -34.0987 | -11.1992
6.2141 | -34.0991 | -11.21
6.2202 | -34.0995 | -11.221
6.2263 | -34.0999 | -11.2315
6.2325 | -34.1002 | -11.2415
6.2386 | -34.1006 | -11.2518
6.2447 | -34.101 | -11.2626
6.2508 | -34.1013 | -11.2731
6.2569 | -34.1017 | -11.284
6.263 | -34.1021 | -11.2942
6.2691 | -34.1025 | -11.3047
6.2752 | -34.1028 | -11.3152
6.2813 | -34.1032 | -11.3257
6.2874 | -34.1036 | -11.3363
6.2935 | -34.104 | -11.3466
6.2996 | -34.1043 | -11.3573
6.3057 | -34.1047 | -11.3677
6.3118 | -34.1051 | -11.3782
6.3179 | -34.1055 | -11.3887
6.324 | -34.1058 | -11.3995
6.3301 | -34.1062 | -11.4098
6.3362 | -34.1066 | -11.4204
6.3423 | -34.1071 | -11.431
6.3484 | -34.1075 | -11.4417
6.3545 | -34.1079 | -11.4517
6.3606 | -34.1083 | -11.4619
6.3667 | -34.1087 | -11.4722
6.3728 | -34.1091 | -11.4828
6.379 | -34.1095 | -11.4933
6.3851 | -34.1098 | -11.5039
6.3912 | -34.1102 | -11.5141
6.3973 | -34.1107 | -11.5243
6.4034 | -34.111 | -11.5348
6.4095 | -34.1114 | -11.5455
6.4156 | -34.1118 | -11.5557
6.4217 | -34.1121 | -11.5662
6.4278 | -34.1125 | -11.5765
6.4339 | -34.113 | -11.587
6.44 | -34.1134 | -11.5971
6.4461 | -34.1137 | -11.608
6.4522 | -34.1141 | -11.6179
6.4583 | -34.1144 | -11.6284
6.4644 | -34.1147 | -11.6392
6.4705 | -34.1152 | -11.6498
6.4766 | -34.1156 | -11.6607
6.4827 | -34.116 | -11.6717
6.4888 | -34.1163 | -11.682
6.4949 | -34.1167 | -11.692
6.501 | -34.1171 | -11.7028
6.5071 | -34.1175 | -11.7134
6.5132 | -34.1179 | -11.7237
6.5194 | -34.1183 | -11.7341
6.5255 | -34.1187 | -11.7442
6.5316 | -34.1191 | -11.7548
6.5377 | -34.1195 | -11.7655
6.5438 | -34.1199 | -11.7758
6.5499 | -34.1203 | -11.7861
6.556 | -34.1207 | -11.7964
6.5621 | -34.121 | -11.807
6.5682 | -34.1214 | -11.8173
6.5743 | -34.1218 | -11.8285
6.5804 | -34.1221 | -11.8389
6.5865 | -34.1224 | -11.8496
6.5926 | -34.1229 | -11.8601
6.5987 | -34.1232 | -11.8702
6.6048 | -34.1236 | -11.8813
6.6109 | -34.124 | -11.8917
6.617 | -34.1244 | -11.903
6.6231 | -34.1247 | -11.9137
6.6292 | -34.1252 | -11.9241
6.6353 | -34.1255 | -11.9355
6.6414 | -34.126 | -11.9463
6.6475 | -34.1264 | -11.957
6.6536 | -34.1268 | -11.9671
6.6597 | -34.1272 | -11.9773
6.6659 | -34.1276 | -11.9876
6.672 | -34.128 | -11.9981
6.6781 | -34.1284 | -12.0084
6.6842 | -34.1288 | -12.0191
6.6903 | -34.1292 | -12.0295
6.6964 | -34.1296 | -12.04
6.7025 | -34.1299 | -12.0505
6.7086 | -34.1303 | -12.0608
6.7147 | -34.1307 | -12.0714
6.7208 | -34.1311 | -12.0824
6.7269 | -34.1314 | -12.0925
6.733 | -34.1317 | -12.1022
6.7391 | -34.1322 | -12.1129
6.7452 | -34.1327 | -12.1228
6.7513 | -34.1331 | -12.1329
6.7574 | -34.1335 | -12.1436
6.7635 | -34.1339 | -12.1536
6.7696 | -34.1344 | -12.1641
6.7757 | -34.1348 | -12.1749
6.7818 | -34.1352 | -12.1854
6.7879 | -34.1356 | -12.1959
6.794 | -34.136 | -12.2062
6.8001 | -34.1364 | -12.2165
6.8063 | -34.1368 | -12.2271
6.8124 | -34.1371 | -12.2376
6.8185 | -34.1375 | -12.248
6.8246 | -34.1379 | -12.2585
6.8307 | -34.1383 | -12.2688
6.8368 | -34.1387 | -12.2795
6.8429 | -34.1391 | -12.2899
6.849 | -34.1395 | -12.3003
6.8551 | -34.1399 | -12.3108
6.8612 | -34.1403 | -12.321
6.8673 | -34.1407 | -12.3312
6.8734 | -34.1411 | -12.3424
6.8795 | -34.1416 | -12.3531
6.8856 | -34.1419 | -12.363
6.8917 | -34.1423 | -12.3729
6.8978 | -34.1427 | -12.3833
6.9039 | -34.1431 | -12.3935
6.91 | -34.1435 | -12.404
6.9161 | -34.1439 | -12.4142
6.9222 | -34.1443 | -12.4247
6.9283 | -34.1446 | -12.4347
6.9344 | -34.145 | -12.4449
6.9405 | -34.1454 | -12.4551
6.9466 | -34.1458 | -12.4657
6.9528 | -34.1462 | -12.4761
6.9589 | -34.1466 | -12.4863
6.965 | -34.147 | -12.4967
6.9711 | -34.1474 | -12.5071
6.9772 | -34.1479 | -12.5172
6.9833 | -34.1482 | -12.5278
6.9894 | -34.1486 | -12.5386
6.9955 | -34.1489 | -12.5489
7.0016 | -34.1493 | -12.5596
7.0077 | -34.1497 | -12.5703
7.0138 | -34.1501 | -12.581
7.0199 | -34.1505 | -12.5911
7.026 | -34.1508 | -12.6014
7.0321 | -34.1512 | -12.6118
7.0382 | -34.1515 | -12.6219
7.0443 | -34.152 | -12.6325
7.0504 | -34.1523 | -12.6432
7.0565 | -34.1527 | -12.6538
7.0626 | -34.1531 | -12.664
7.0687 | -34.1534 | -12.674
7.0748 | -34.1538 | -12.6842
7.0809 | -34.1542 | -12.6946
7.087 | -34.1546 | -12.7051
7.0932 | -34.1549 | -12.7153
7.0993 | -34.1553 | -12.7258
7.1054 | -34.1556 | -12.7358
7.1115 | -34.156 | -12.7461
7.1176 | -34.1563 | -12.7562
7.1237 | -34.1566 | -12.7666
7.1298 | -34.1569 | -12.7764
7.1359 | -34.1573 | -12.7864
7.142 | -34.1576 | -12.7968
7.1481 | -34.158 | -12.8068
7.1542 | -34.1583 | -12.817
7.1603 | -34.1587 | -12.8271
7.1664 | -34.1591 | -12.8375
7.1725 | -34.1595 | -12.8475
7.1786 | -34.1599 | -12.8579
7.1847 | -34.1602 | -12.868
7.1908 | -34.1606 | -12.8784
7.1969 | -34.1609 | -12.8886
7.203 | -34.1613 | -12.8988
7.2091 | -34.1616 | -12.9088
7.2152 | -34.162 | -12.9191
7.2213 | -34.1623 | -12.9291
7.2274 | -34.1627 | -12.9393
7.2335 | -34.163 | -12.9495
7.2397 | -34.1634 | -12.9597
7.2458 | -34.1638 | -12.9694
7.2519 | -34.1641 | -12.9796
7.258 | -34.1645 | -12.99
7.2641 | -34.1648 | -13.0001
7.2702 | -34.1652 | -13.0106
7.2763 | -34.1656 | -13.0205
7.2824 | -34.166 | -13.0306
7.2885 | -34.1664 | -13.0407
7.2946 | -34.1668 | -13.0511
7.3007 | -34.167 | -13.0613
7.3068 | -34.1674 | -13.0717
7.3129 | -34.1677 | -13.0816
7.319 | -34.1681 | -13.0919
7.3251 | -34.1685 | -13.1017
7.3312 | -34.1689 | -13.1115
7.3373 | -34.1693 | -13.1213
7.3434 | -34.1697 | -13.1312
7.3495 | -34.1701 | -13.1409
7.3556 | -34.1704 | -13.1514
7.3617 | -34.1707 | -13.162
7.3678 | -34.1711 | -13.1722
7.3739 | -34.1715 | -13.1819
7.3801 | -34.1718 | -13.1919
7.3862 | -34.1722 | -13.2019
7.3923 | -34.1725 | -13.212
7.3984 | -34.1729 | -13.2219
7.4045 | -34.1732 | -13.2319
7.4106 | -34.1736 | -13.2418
7.4167 | -34.1739 | -13.2521
7.4228 | -34.1744 | -13.2621
7.4289 | -34.1747 | -13.2718
7.435 | -34.1751 | -13.2816
7.4411 | -34.1755 | -13.2914
7.4472 | -34.1758 | -13.3016
7.4533 | -34.1762 | -13.3116
7.4594 | -34.1766 | -13.322
7.4655 | -34.177 | -13.332
7.4716 | -34.1773 | -13.3423
7.4777 | -34.1777 | -13.3518
7.4838 | -34.178 | -13.362
7.4899 | -34.1784 | -13.3721
7.496 | -34.1788 | -13.3821
7.5021 | -34.1792 | -13.392
7.5082 | -34.1795 | -13.4021
7.5143 | -34.1799 | -13.412
7.5204 | -34.1802 | -13.4217
7.5266 | -34.1806 | -13.4318
7.5327 | -34.1809 | -13.4417
7.5388 | -34.1812 | -13.4516
7.5449 | -34.1816 | -13.4619
7.551 | -34.1819 | -13.4721
7.5571 | -34.1823 | -13.482
7.5632 | -34.1826 | -13.4919
7.5693 | -34.183 | -13.5021
7.5754 | -34.1834 | -13.5121
7.5815 | -34.1838 | -13.5217
7.5876 | -34.1841 | -13.5317
7.5937 | -34.1844 | -13.5416
7.5998 | -34.1848 | -13.5516
7.6059 | -34.1851 | -13.5616
7.612 | -34.1855 | -13.5715
7.6181 | -34.1858 | -13.5814
7.6242 | -34.1862 | -13.5914
7.6303 | -34.1866 | -13.6012
7.6364 | -34.187 | -13.6114
7.6425 | -34.1873 | -13.6217
7.6486 | -34.1877 | -13.6317
7.6547 | -34.188 | -13.6418
7.6608 | -34.1884 | -13.6518
7.667 | -34.1888 | -13.6612
7.6731 | -34.1891 | -13.6713
7.6792 | -34.1894 | -13.6815
7.6853 | -34.1897 | -13.6913
7.6914 | -34.19 | -13.7014
7.6975 | -34.1904 | -13.7116
7.7036 | -34.1907 | -13.7217
7.7097 | -34.1911 | -13.7316
7.7158 | -34.1914 | -13.7414
7.7219 | -34.1917 | -13.7513
7.728 | -34.192 | -13.7614
7.7341 | -34.1924 | -13.7714
7.7402 | -34.1928 | -13.7813
7.7463 | -34.1931 | -13.7913
7.7524 | -34.1934 | -13.8016
7.7585 | -34.1937 | -13.8117
7.7646 | -34.1941 | -13.8214
7.7707 | -34.1943 | -13.8311
7.7768 | -34.1946 | -13.8409
7.7829 | -34.1949 | -13.851
7.789 | -34.1952 | -13.8607
7.7951 | -34.1956 | -13.8711
7.8012 | -34.1959 | -13.8814
7.8073 | -34.1962 | -13.8913
7.8135 | -34.1966 | -13.9013
7.8196 | -34.1969 | -13.9112
7.8257 | -34.1973 | -13.9215
7.8318 | -34.1976 | -13.9319
7.8379 | -34.198 | -13.9414
7.844 | -34.1982 | -13.951
7.8501 | -34.1985 | -13.9613
7.8562 | -34.1988 | -13.9719
7.8623 | -34.1992 | -13.9819
7.8684 | -34.1995 | -13.9916
7.8745 | -34.1998 | -14.0014
7.8806 | -34.2001 | -14.0109
7.8867 | -34.2005 | -14.0207
7.8928 | -34.2008 | -14.0306
7.8989 | -34.2012 | -14.04
7.905 | -34.2016 | -14.0504
7.9111 | -34.2019 | -14.0604
7.9172 | -34.2022 | -14.0701
7.9233 | -34.2025 | -14.0801
7.9294 | -34.2029 | -14.0902
7.9355 | -34.2032 | -14.0997
7.9416 | -34.2034 | -14.1096
7.9477 | -34.2038 | -14.1191
7.9539 | -34.2041 | -14.1286
7.96 | -34.2045 | -14.1388
7.9661 | -34.2048 | -14.1489
7.9722 | -34.2052 | -14.1589
7.9783 | -34.2055 | -14.1688
7.9844 | -34.2059 | -14.1791
7.9905 | -34.2061 | -14.1889
7.9966 | -34.2064 | -14.1989
8.0027 | -34.2068 | -14.2091
8.0088 | -34.2071 | -14.2187
8.0149 | -34.2074 | -14.2283
8.021 | -34.2077 | -14.2383
8.0271 | -34.2081 | -14.2487
8.0332 | -34.2084 | -14.2586
8.0393 | -34.2087 | -14.2687
8.0454 | -34.209 | -14.2788
8.0515 | -34.2094 | -14.2887
8.0576 | -34.2096 | -14.2989
8.0637 | -34.21 | -14.3083
8.0698 | -34.2103 | -14.3179
8.0759 | -34.2106 | -14.3275
8.082 | -34.211 | -14.3376
8.0881 | -34.2113 | -14.3474
8.0942 | -34.2116 | -14.3578
8.1004 | -34.2119 | -14.368
8.1065 | -34.2122 | -14.3779
8.1126 | -34.2126 | -14.3878
8.1187 | -34.213 | -14.3978
8.1248 | -34.2133 | -14.408
8.1309 | -34.2136 | -14.4179
8.137 | -34.2139 | -14.4277
8.1431 | -34.2143 | -14.4378
8.1492 | -34.2146 | -14.4478
8.1553 | -34.2149 | -14.4578
8.1614 | -34.2153 | -14.468
8.1675 | -34.2155 | -14.4777
8.1736 | -34.2159 | -14.488
8.1797 | -34.2162 | -14.4981
8.1858 | -34.2165 | -14.508
8.1919 | -34.2168 | -14.5178
8.198 | -34.2171 | -14.5274
8.2041 | -34.2175 | -14.5377
8.2102 | -34.2178 | -14.5476
8.2163 | -34.2182 | -14.5575
8.2224 | -34.2185 | -14.5677
8.2285 | -34.2188 | -14.5775
8.2346 | -34.2192 | -14.5874
8.2408 | -34.2195 | -14.5974
8.2469 | -34.2198 | -14.6073
8.253 | -34.2201 | -14.6171
8.2591 | -34.2205 | -14.6268
8.2652 | -34.2209 | -14.6365
8.2713 | -34.2212 | -14.6465
8.2774 | -34.2215 | -14.6565
8.2835 | -34.2219 | -14.6669
8.2896 | -34.2222 | -14.6763
8.2957 | -34.2226 | -14.6863
8.3018 | -34.2229 | -14.6962
8.3079 | -34.2233 | -14.7061
8.314 | -34.2236 | -14.7158
8.3201 | -34.2239 | -14.7257
8.3262 | -34.2242 | -14.7355
8.3323 | -34.2246 | -14.7454
8.3384 | -34.2249 | -14.7558
8.3445 | -34.2253 | -14.7653
8.3506 | -34.2256 | -14.7755
8.3567 | -34.226 | -14.7856
8.3628 | -34.2263 | -14.7955
8.3689 | -34.2266 | -14.8054
8.375 | -34.2269 | -14.8154
8.3812 | -34.2273 | -14.8252
8.3873 | -34.2276 | -14.8353
8.3934 | -34.2279 | -14.8452
8.3995 | -34.2283 | -14.8552
8.4056 | -34.2286 | -14.8652
8.4117 | -34.2289 | -14.8754
8.4178 | -34.2293 | -14.8849
8.4239 | -34.2296 | -14.8946
8.43 | -34.23 | -14.9045
8.4361 | -34.2303 | -14.9148
8.4422 | -34.2307 | -14.9247
8.4483 | -34.2309 | -14.9346
8.4544 | -34.2313 | -14.9448
8.4605 | -34.2316 | -14.9547
8.4666 | -34.2319 | -14.9648
8.4727 | -34.2322 | -14.9746
8.4788 | -34.2326 | -14.9844
8.4849 | -34.2329 | -14.9946
8.491 | -34.2332 | -15.0044
8.4971 | -34.2336 | -15.0144
8.5032 | -34.2339 | -15.0244
8.5093 | -34.2343 | -15.0343
8.5154 | -34.2346 | -15.044
8.5215 | -34.2349 | -15.054
8.5277 | -34.2352 | -15.064
8.5338 | -34.2356 | -15.0737
8.5399 | -34.2359 | -15.0839
8.546 | -34.2363 | -15.0935
8.5521 | -34.2366 | -15.1036
8.5582 | -34.2369 | -15.1137
8.5643 | -34.2373 | -15.1234
8.5704 | -34.2376 | -15.133
8.5765 | -34.2379 | -15.1428
8.5826 | -34.2384 | -15.1527
8.5887 | -34.2387 | -15.1626
8.5948 | -34.239 | -15.1721
8.6009 | -34.2394 | -15.1819
8.607 | -34.2397 | -15.192
8.6131 | -34.24 | -15.2022
8.6192 | -34.2403 | -15.2118
8.6253 | -34.2407 | -15.2218
8.6314 | -34.2411 | -15.2317
8.6375 | -34.2414 | -15.2418
8.6436 | -34.2418 | -15.2518
8.6497 | -34.2421 | -15.2618
8.6558 | -34.2425 | -15.2714
8.6619 | -34.2428 | -15.2815
8.6681 | -34.2431 | -15.2913
8.6742 | -34.2434 | -15.3014
8.6803 | -34.2438 | -15.3116
8.6864 | -34.2442 | -15.3218
8.6925 | -34.2446 | -15.3317
8.6986 | -34.2449 | -15.342
8.7047 | -34.2453 | -15.352
8.7108 | -34.2456 | -15.362
8.7169 | -34.2459 | -15.3716
8.723 | -34.2463 | -15.3816
8.7291 | -34.2466 | -15.3914
8.7352 | -34.247 | -15.4016
8.7413 | -34.2474 | -15.4114
8.7474 | -34.2477 | -15.421
8.7535 | -34.2481 | -15.4311
8.7596 | -34.2485 | -15.4415
8.7657 | -34.2488 | -15.4514
8.7718 | -34.2492 | -15.461
8.7779 | -34.2495 | -15.4714
8.784 | -34.2498 | -15.4812
8.7901 | -34.2501 | -15.4915
8.7962 | -34.2505 | -15.5016
8.8023 | -34.2509 | -15.5111
8.8084 | -34.2513 | -15.5213
8.8146 | -34.2516 | -15.5311
8.8207 | -34.252 | -15.5411
8.8268 | -34.2524 | -15.5509
8.8329 | -34.2527 | -15.561
8.839 | -34.2531 | -15.571
8.8451 | -34.2534 | -15.5807
8.8512 | -34.2539 | -15.5907
8.8573 | -34.2542 | -15.6005
8.8634 | -34.2546 | -15.6103
8.8695 | -34.2549 | -15.6204
8.8756 | -34.2553 | -15.6301
8.8817 | -34.2557 | -15.6403
8.8878 | -34.2559 | -15.6504
8.8939 | -34.2563 | -15.6601
8.9 | -34.2566 | -15.6696
8.9061 | -34.257 | -15.6796
8.9122 | -34.2574 | -15.6894
8.9183 | -34.2577 | -15.6992
8.9244 | -34.2581 | -15.709
8.9305 | -34.2585 | -15.719
8.9366 | -34.2589 | -15.7287
8.9427 | -34.2593 | -15.7387
8.9488 | -34.2596 | -15.7485
8.955 | -34.26 | -15.7584
8.9611 | -34.2604 | -15.7684
8.9672 | -34.2608 | -15.7788
8.9733 | -34.2611 | -15.7885
8.9794 | -34.2615 | -15.7986
8.9855 | -34.2619 | -15.8084
8.9916 | -34.2622 | -15.8182
8.9977 | -34.2626 | -15.828
9.0038 | -34.263 | -15.8381
9.0099 | -34.2634 | -15.8481
9.016 | -34.2638 | -15.858
9.0221 | -34.2642 | -15.8681
9.0282 | -34.2646 | -15.8782
9.0343 | -34.265 | -15.8884
9.0404 | -34.2654 | -15.8981
9.0465 | -34.2658 | -15.9084
9.0526 | -34.2662 | -15.9183
9.0587 | -34.2666 | -15.9281
9.0648 | -34.267 | -15.938
9.0709 | -34.2674 | -15.9477
9.077 | -34.2677 | -15.9582
9.0831 | -34.2681 | -15.9681
9.0892 | -34.2684 | -15.978
9.0953 | -34.2688 | -15.988
9.1015 | -34.2691 | -15.9979
9.1076 | -34.2696 | -16.0079
9.1137 | -34.27 | -16.0178
9.1198 | -34.2704 | -16.0273
9.1259 | -34.2708 | -16.0372
9.132 | -34.2712 | -16.0474
9.1381 | -34.2716 | -16.0572
9.1442 | -34.2721 | -16.0667
9.1503 | -34.2724 | -16.0767
9.1564 | -34.2728 | -16.0867
9.1625 | -34.2731 | -16.0967
9.1686 | -34.2736 | -16.1066
9.1747 | -34.274 | -16.1166
9.1808 | -34.2744 | -16.1268
9.1869 | -34.2747 | -16.1369
9.193 | -34.2751 | -16.147
9.1991 | -34.2755 | -16.1567
9.2052 | -34.2759 | -16.1665
9.2113 | -34.2763 | -16.1767
9.2174 | -34.2767 | -16.1872
9.2235 | -34.2771 | -16.1969
9.2296 | -34.2776 | -16.207
9.2357 | -34.278 | -16.2169
9.2419 | -34.2783 | -16.2268
9.248 | -34.2788 | -16.237
9.2541 | -34.2792 | -16.247
9.2602 | -34.2795 | -16.2571
9.2663 | -34.28 | -16.2668
9.2724 | -34.2803 | -16.2769
9.2785 | -34.2807 | -16.2868
9.2846 | -34.2812 | -16.2967
9.2907 | -34.2815 | -16.3066
9.2968 | -34.2819 | -16.3167
9.3029 | -34.2823 | -16.3268
9.309 | -34.2827 | -16.3366
9.3151 | -34.2831 | -16.3464
9.3212 | -34.2835 | -16.3565
9.3273 | -34.284 | -16.3664
9.3334 | -34.2844 | -16.3761
9.3395 | -34.2849 | -16.3859
9.3456 | -34.2853 | -16.3956
9.3517 | -34.2857 | -16.4051
9.3578 | -34.2861 | -16.4148
9.3639 | -34.2866 | -16.4249
9.37 | -34.287 | -16.4349
9.3761 | -34.2874 | -16.4447
9.3822 | -34.2878 | -16.4546
9.3884 | -34.2882 | -16.4644
9.3945 | -34.2886 | -16.4744
9.4006 | -34.289 | -16.4843
9.4067 | -34.2894 | -16.4944
9.4128 | -34.2899 | -16.5043
9.4189 | -34.2903 | -16.5143
9.425 | -34.2907 | -16.5242
9.4311 | -34.2911 | -16.534
9.4372 | -34.2915 | -16.544
9.4433 | -34.2919 | -16.554
9.4494 | -34.2923 | -16.5638
9.4555 | -34.2928 | -16.5737
9.4616 | -34.2932 | -16.5834
9.4677 | -34.2937 | -16.5932
9.4738 | -34.2941 | -16.6034
9.4799 | -34.2945 | -16.6132
9.486 | -34.2948 | -16.6234
9.4921 | -34.2953 | -16.6338
9.4982 | -34.2957 | -16.6432
9.5043 | -34.2961 | -16.6528
9.5104 | -34.2965 | -16.6626
9.5165 | -34.297 | -16.6725
9.5226 | -34.2974 | -16.6826
9.5288 | -34.2978 | -16.6924
9.5349 | -34.2983 | -16.7024
9.541 | -34.2987 | -16.7123
9.5471 | -34.2991 | -16.7223
9.5532 | -34.2995 | -16.7322
9.5593 | -34.3 | -16.742
9.5654 | -34.3004 | -16.752
9.5715 | -34.3008 | -16.7621
9.5776 | -34.3012 | -16.7723
9.5837 | -34.3016 | -16.7819
9.5898 | -34.302 | -16.792
9.5959 | -34.3024 | -16.8018
9.602 | -34.3028 | -16.8117
9.6081 | -34.3032 | -16.8217
9.6142 | -34.3037 | -16.8315
9.6203 | -34.3041 | -16.8414
9.6264 | -34.3044 | -16.8509
9.6325 | -34.3048 | -16.861
9.6386 | -34.3052 | -16.8708
9.6447 | -34.3056 | -16.8806
9.6508 | -34.3061 | -16.8906
9.6569 | -34.3065 | -16.9005
9.663 | -34.3069 | -16.9105
9.6691 | -34.3074 | -16.9204
9.6753 | -34.3079 | -16.9303
9.6814 | -34.3083 | -16.94
9.6875 | -34.3088 | -16.9497
9.6936 | -34.3092 | -16.9594
9.6997 | -34.3096 | -16.9695
9.7058 | -34.3101 | -16.9795
9.7119 | -34.3105 | -16.9894
9.718 | -34.3109 | -16.9995
9.7241 | -34.3113 | -17.0095
9.7302 | -34.3117 | -17.0194
9.7363 | -34.3121 | -17.0291
9.7424 | -34.3126 | -17.0387
9.7485 | -34.313 | -17.0485
9.7546 | -34.3135 | -17.0582
9.7607 | -34.3138 | -17.0684
9.7668 | -34.3143 | -17.0786
9.7729 | -34.3147 | -17.0882
9.779 | -34.3153 | -17.0982
9.7851 | -34.3156 | -17.1082
9.7912 | -34.316 | -17.1186
9.7973 | -34.3164 | -17.1287
9.8034 | -34.3169 | -17.1388
9.8095 | -34.3173 | -17.1486
9.8157 | -34.3177 | -17.1585
9.8218 | -34.3181 | -17.168
9.8279 | -34.3186 | -17.178
9.834 | -34.319 | -17.1877
9.8401 | -34.3195 | -17.1975
9.8462 | -34.32 | -17.207
9.8523 | -34.3204 | -17.2169
9.8584 | -34.3208 | -17.2269
9.8645 | -34.3212 | -17.2369
9.8706 | -34.3216 | -17.2466
9.8767 | -34.322 | -17.2561
9.8828 | -34.3225 | -17.2654
9.8889 | -34.3229 | -17.2756
9.895 | -34.3233 | -17.2851
9.9011 | -34.3237 | -17.2945
9.9072 | -34.3242 | -17.304
9.9133 | -34.3246 | -17.3137
9.9194 | -34.3251 | -17.3233
9.9255 | -34.3255 | -17.3333
9.9316 | -34.326 | -17.3428
9.9377 | -34.3264 | -17.3525
9.9438 | -34.3269 | -17.3623
9.9499 | -34.3273 | -17.3721
9.956 | -34.3277 | -17.3817
9.9622 | -34.3282 | -17.3914
9.9683 | -34.3286 | -17.4013
9.9744 | -34.329 | -17.411
9.9805 | -34.3295 | -17.4209
9.9866 | -34.3299 | -17.4304
9.9927 | -34.3304 | -17.4404
9.9988 | -34.3308 | -17.45
10.0049 | -34.3312 | -17.4598
10.011 | -34.3317 | -17.4695
10.0171 | -34.3321 | -17.4793
10.0232 | -34.3326 | -17.4888
10.0293 | -34.3331 | -17.4983
10.0354 | -34.3336 | -17.5078
10.0415 | -34.334 | -17.5175
10.0476 | -34.3344 | -17.5276
10.0537 | -34.3348 | -17.5376
10.0598 | -34.3353 | -17.5475
10.0659 | -34.3357 | -17.5573
10.072 | -34.3362 | -17.5669
10.0781 | -34.3367 | -17.5767
10.0842 | -34.3371 | -17.5865
10.0903 | -34.3375 | -17.5965
10.0964 | -34.3381 | -17.606
10.1026 | -34.3385 | -17.6156
10.1087 | -34.339 | -17.6254
10.1148 | -34.3394 | -17.6353
10.1209 | -34.3398 | -17.6448
10.127 | -34.3403 | -17.6544
10.1331 | -34.3408 | -17.664
10.1392 | -34.3412 | -17.6736
10.1453 | -34.3417 | -17.6834
10.1514 | -34.3421 | -17.6929
10.1575 | -34.3425 | -17.7025
10.1636 | -34.343 | -17.7114
10.1697 | -34.3435 | -17.7211
10.1758 | -34.3439 | -17.7306
10.1819 | -34.3444 | -17.7404
10.188 | -34.3448 | -17.7498
10.1941 | -34.3452 | -17.7596
10.2002 | -34.3456 | -17.7686
10.2063 | -34.3461 | -17.7785
10.2124 | -34.3465 | -17.7882
10.2185 | -34.347 | -17.7977
10.2246 | -34.3474 | -17.8073
10.2307 | -34.3478 | -17.8171
10.2368 | -34.3483 | -17.8258
10.2429 | -34.3488 | -17.8352
10.2491 | -34.3492 | -17.8446
10.2552 | -34.3497 | -17.8547
10.2613 | -34.3502 | -17.8645
10.2674 | -34.3506 | -17.8743
10.2735 | -34.351 | -17.8841
10.2796 | -34.3514 | -17.8939
10.2857 | -34.3518 | -17.9036
10.2918 | -34.3523 | -17.9134
10.2979 | -34.3528 | -17.9231
10.304 | -34.3533 | -17.9331
10.3101 | -34.3537 | -17.9426
10.3162 | -34.3542 | -17.9522
10.3223 | -34.3546 | -17.9618
10.3284 | -34.3551 | -17.9714
10.3345 | -34.3555 | -17.9808
10.3406 | -34.3559 | -17.9903
10.3467 | -34.3564 | -17.9998
10.3528 | -34.3568 | -18.0092
10.3589 | -34.3572 | -18.019
10.365 | -34.3577 | -18.0288
10.3711 | -34.3581 | -18.0387
10.3772 | -34.3585 | -18.0484
10.3833 | -34.3589 | -18.058
10.3895 | -34.3594 | -18.0673
10.3956 | -34.3598 | -18.0767
10.4017 | -34.3602 | -18.0862
10.4078 | -34.3607 | -18.0958
10.4139 | -34.3612 | -18.1055
10.42 | -34.3616 | -18.1151
10.4261 | -34.3621 | -18.1248
10.4322 | -34.3626 | -18.1341
10.4383 | -34.363 | -18.1438
10.4444 | -34.3634 | -18.1532
10.4505 | -34.3639 | -18.1627
10.4566 | -34.3643 | -18.1722
10.4627 | -34.3648 | -18.1819
10.4688 | -34.3653 | -18.1911
10.4749 | -34.3657 | -18.2006
10.481 | -34.3662 | -18.2103
10.4871 | -34.3666 | -18.2199
10.4932 | -34.3671 | -18.2294
10.4993 | -34.3676 | -18.2389
10.5054 | -34.368 | -18.2484
10.5115 | -34.3685 | -18.2578
10.5176 | -34.3689 | -18.2673
10.5237 | -34.3694 | -18.2768
10.5298 | -34.3698 | -18.2861
10.536 | -34.3704 | -18.2958
10.5421 | -34.3708 | -18.3049
10.5482 | -34.3712 | -18.3147
10.5543 | -34.3717 | -18.3243
10.5604 | -34.3722 | -18.3334
10.5665 | -34.3726 | -18.3427
10.5726 | -34.373 | -18.3519
10.5787 | -34.3735 | -18.3615
10.5848 | -34.374 | -18.3706
10.5909 | -34.3745 | -18.3801
10.597 | -34.3749 | -18.3896
10.6031 | -34.3754 | -18.3993
10.6092 | -34.3758 | -18.409
10.6153 | -34.3762 | -18.4188
10.6214 | -34.3767 | -18.4282
10.6275 | -34.3772 | -18.4376
10.6336 | -34.3777 | -18.4474
10.6397 | -34.3781 | -18.4567
10.6458 | -34.3785 | -18.4658
10.6519 | -34.3789 | -18.4752
10.658 | -34.3793 | -18.4845
10.6641 | -34.3798 | -18.494
10.6702 | -34.3803 | -18.5037
10.6764 | -34.3808 | -18.513
10.6825 | -34.3813 | -18.5227
10.6886 | -34.3818 | -18.5322
10.6947 | -34.3822 | -18.5416
10.7008 | -34.3827 | -18.551
10.7069 | -34.3832 | -18.5606
10.713 | -34.3837 | -18.5699
10.7191 | -34.3842 | -18.5795
10.7252 | -34.3847 | -18.5886
10.7313 | -34.3852 | -18.5981
10.7374 | -34.3856 | -18.6074
10.7435 | -34.386 | -18.617
10.7496 | -34.3865 | -18.626
10.7557 | -34.387 | -18.6355
10.7618 | -34.3875 | -18.645
10.7679 | -34.388 | -18.6544
10.774 | -34.3884 | -18.6638
10.7801 | -34.3888 | -18.6736
10.7862 | -34.3893 | -18.683
10.7923 | -34.3898 | -18.6927
10.7984 | -34.3903 | -18.7024
10.8045 | -34.3907 | -18.712
10.8106 | -34.3912 | -18.7213
10.8168 | -34.3917 | -18.7309
10.8229 | -34.3921 | -18.7405
10.829 | -34.3926 | -18.75
10.8351 | -34.3931 | -18.7595
10.8412 | -34.3935 | -18.7687
10.8473 | -34.3939 | -18.7777
10.8534 | -34.3944 | -18.7874
10.8595 | -34.3949 | -18.7968
10.8656 | -34.3954 | -18.8061
10.8717 | -34.3959 | -18.8152
10.8778 | -34.3963 | -18.8245
10.8839 | -34.3968 | -18.8337
10.89 | -34.3972 | -18.8429
10.8961 | -34.3977 | -18.8524
10.9022 | -34.3982 | -18.8616
10.9083 | -34.3987 | -18.871
10.9144 | -34.3991 | -18.8803
10.9205 | -34.3996 | -18.8895
10.9266 | -34.4001 | -18.8993
10.9327 | -34.4007 | -18.9086
10.9388 | -34.4011 | -18.918
10.9449 | -34.4016 | -18.9274
10.951 | -34.4021 | -18.9366
10.9571 | -34.4025 | -18.9461
10.9633 | -34.403 | -18.9558
10.9694 | -34.4035 | -18.9657
10.9755 | -34.404 | -18.9751
10.9816 | -34.4044 | -18.9845
10.9877 | -34.4049 | -18.9938
10.9938 | -34.4053 | -19.0033
10.9999 | -34.4058 | -19.0128
11.006 | -34.4063 | -19.0221
11.0121 | -34.4067 | -19.0317
11.0182 | -34.4072 | -19.0407
11.0243 | -34.4077 | -19.05
11.0304 | -34.4081 | -19.0592
11.0365 | -34.4086 | -19.0683
11.0426 | -34.4091 | -19.0774
11.0487 | -34.4096 | -19.0869
11.0548 | -34.4101 | -19.096
11.0609 | -34.4106 | -19.1059
11.067 | -34.411 | -19.1154
11.0731 | -34.4115 | -19.1246
11.0792 | -34.4121 | -19.1339
11.0853 | -34.4126 | -19.143
11.0914 | -34.4131 | -19.1521
11.0975 | -34.4136 | -19.1618
11.1037 | -34.414 | -19.1708
11.1098 | -34.4145 | -19.1805
11.1159 | -34.415 | -19.1899
11.122 | -34.4155 | -19.1997
11.1281 | -34.4159 | -19.209
11.1342 | -34.4164 | -19.218
11.1403 | -34.4169 | -19.2276
11.1464 | -34.4173 | -19.2369
11.1525 | -34.4178 | -19.2464
11.1586 | -34.4182 | -19.2564
11.1647 | -34.4187 | -19.2658
11.1708 | -34.4192 | -19.2747
11.1769 | -34.4196 | -19.2836
11.183 | -34.4201 | -19.293
11.1891 | -34.4206 | -19.3027
11.1952 | -34.421 | -19.3123
11.2013 | -34.4215 | -19.3215
11.2074 | -34.4219 | -19.3305
11.2135 | -34.4224 | -19.34
11.2196 | -34.4228 | -19.3492
11.2257 | -34.4233 | -19.3584
11.2318 | -34.4237 | -19.3679
11.2379 | -34.4241 | -19.3772
11.244 | -34.4246 | -19.3864
11.2502 | -34.4251 | -19.3958
11.2563 | -34.4256 | -19.4048
11.2624 | -34.426 | -19.414
11.2685 | -34.4265 | -19.4232
11.2746 | -34.4269 | -19.4323
11.2807 | -34.4274 | -19.4422
11.2868 | -34.4279 | -19.4518
11.2929 | -34.4284 | -19.4608
11.299 | -34.4289 | -19.4702
11.3051 | -34.4293 | -19.4794
11.3112 | -34.4297 | -19.4888
11.3173 | -34.4303 | -19.4982
11.3234 | -34.4307 | -19.5076
11.3295 | -34.4311 | -19.5167
11.3356 | -34.4316 | -19.5262
11.3417 | -34.4321 | -19.5359
11.3478 | -34.4325 | -19.5449
11.3539 | -34.4329 | -19.5542
11.36 | -34.4334 | -19.5638
11.3661 | -34.4339 | -19.5729
11.3722 | -34.4343 | -19.5821
11.3783 | -34.4348 | -19.5914
11.3844 | -34.4352 | -19.6008
11.3906 | -34.4356 | -19.6102
11.3967 | -34.4361 | -19.6192
11.4028 | -34.4366 | -19.6287
11.4089 | -34.4371 | -19.6377
11.415 | -34.4376 | -19.6471
11.4211 | -34.438 | -19.6562
11.4272 | -34.4384 | -19.6656
11.4333 | -34.4389 | -19.6752
11.4394 | -34.4394 | -19.6848
11.4455 | -34.4398 | -19.694
11.4516 | -34.4402 | -19.7031
11.4577 | -34.4407 | -19.7126
11.4638 | -34.4412 | -19.7217
11.4699 | -34.4416 | -19.7315
11.476 | -34.4421 | -19.741
11.4821 | -34.4426 | -19.7501
11.4882 | -34.4431 | -19.7596
11.4943 | -34.4436 | -19.7688
11.5004 | -34.4441 | -19.7782
11.5065 | -34.4446 | -19.7875
11.5126 | -34.4451 | -19.7963
11.5187 | -34.4456 | -19.8056
11.5248 | -34.446 | -19.8149
11.5309 | -34.4465 | -19.824
11.5371 | -34.447 | -19.8331
11.5432 | -34.4474 | -19.8426
11.5493 | -34.4478 | -19.852
11.5554 | -34.4483 | -19.8611
11.5615 | -34.4488 | -19.8706
11.5676 | -34.4493 | -19.8801
11.5737 | -34.4498 | -19.8892
11.5798 | -34.4502 | -19.8983
11.5859 | -34.4507 | -19.9075
11.592 | -34.4512 | -19.9169
11.5981 | -34.4517 | -19.926
11.6042 | -34.4522 | -19.9354
11.6103 | -34.4526 | -19.9445
11.6164 | -34.4531 | -19.9534
11.6225 | -34.4536 | -19.9626
11.6286 | -34.4541 | -19.972
11.6347 | -34.4545 | -19.9815
11.6408 | -34.455 | -19.9908
11.6469 | -34.4555 | -20.0002
11.653 | -34.456 | -20.0092
11.6591 | -34.4565 | -20.0185
11.6652 | -34.4569 | -20.0279
11.6713 | -34.4574 | -20.0372
11.6775 | -34.4579 | -20.0463
11.6836 | -34.4584 | -20.0556
11.6897 | -34.4589 | -20.0646
11.6958 | -34.4593 | -20.0743
11.7019 | -34.4598 | -20.0835
11.708 | -34.4603 | -20.0929
11.7141 | -34.4608 | -20.1022
11.7202 | -34.4613 | -20.1111
11.7263 | -34.4617 | -20.1204
11.7324 | -34.4622 | -20.1298
11.7385 | -34.4626 | -20.1391
11.7446 | -34.4631 | -20.1485
11.7507 | -34.4636 | -20.1575
11.7568 | -34.464 | -20.1668
11.7629 | -34.4645 | -20.1762
11.769 | -34.465 | -20.1854
11.7751 | -34.4655 | -20.1943
11.7812 | -34.466 | -20.2033
11.7873 | -34.4664 | -20.2129
11.7934 | -34.467 | -20.2221
11.7995 | -34.4674 | -20.2309
11.8056 | -34.4679 | -20.2406
11.8117 | -34.4683 | -20.2499
11.8178 | -34.4688 | -20.259
11.824 | -34.4694 | -20.2681
11.8301 | -34.4699 | -20.2769
11.8362 | -34.4703 | -20.2858
11.8423 | -34.4708 | -20.295
11.8484 | -34.4713 | -20.3042
11.8545 | -34.4717 | -20.3136
11.8606 | -34.4722 | -20.3228
11.8667 | -34.4727 | -20.3322
11.8728 | -34.4732 | -20.3413
11.8789 | -34.4736 | -20.3509
11.885 | -34.4741 | -20.36
11.8911 | -34.4746 | -20.3693
11.8972 | -34.475 | -20.3782
11.9033 | -34.4755 | -20.387
11.9094 | -34.4759 | -20.3958
11.9155 | -34.4763 | -20.4051
11.9216 | -34.4769 | -20.4141
11.9277 | -34.4774 | -20.4231
11.9338 | -34.4778 | -20.4322
11.9399 | -34.4783 | -20.4417
11.946 | -34.4788 | -20.451
11.9521 | -34.4793 | -20.4599
11.9582 | -34.4798 | -20.4689
11.9644 | -34.4802 | -20.478
11.9705 | -34.4807 | -20.4873
11.9766 | -34.4812 | -20.4963
11.9827 | -34.4816 | -20.5053
11.9888 | -34.4821 | -20.5142
11.9949 | -34.4825 | -20.5234
12.001 | -34.483 | -20.5327
12.0071 | -34.4835 | -20.5415
12.0132 | -34.484 | -20.5504
12.0193 | -34.4845 | -20.5599
12.0254 | -34.4849 | -20.5687
12.0315 | -34.4854 | -20.5781
12.0376 | -34.4859 | -20.5871
12.0437 | -34.4864 | -20.5959
12.0498 | -34.4868 | -20.6052
12.0559 | -34.4873 | -20.6145
12.062 | -34.4878 | -20.6235
12.0681 | -34.4883 | -20.6326
12.0742 | -34.4888 | -20.6418
12.0803 | -34.4893 | -20.6509
12.0864 | -34.4897 | -20.6601
12.0925 | -34.4902 | -20.6692
12.0986 | -34.4908 | -20.6781
12.1047 | -34.4913 | -20.6871
12.1109 | -34.4917 | -20.6963
12.117 | -34.4922 | -20.7053
12.1231 | -34.4927 | -20.7145
12.1292 | -34.4932 | -20.7236
12.1353 | -34.4937 | -20.7324
12.1414 | -34.4942 | -20.7417
12.1475 | -34.4947 | -20.7509
12.1536 | -34.4951 | -20.7601
12.1597 | -34.4956 | -20.7692
12.1658 | -34.4961 | -20.7782
12.1719 | -34.4966 | -20.7872
12.178 | -34.4971 | -20.7966
12.1841 | -34.4975 | -20.8059
12.1902 | -34.4981 | -20.8152
12.1963 | -34.4985 | -20.8241
12.2024 | -34.499 | -20.833
12.2085 | -34.4994 | -20.8422
12.2146 | -34.4999 | -20.8512
12.2207 | -34.5004 | -20.8604
12.2268 | -34.5009 | -20.8696
12.2329 | -34.5014 | -20.8787
12.239 | -34.5018 | -20.8874
12.2451 | -34.5023 | -20.8963
12.2513 | -34.5028 | -20.9054
12.2574 | -34.5033 | -20.9146
12.2635 | -34.5038 | -20.9232
12.2696 | -34.5042 | -20.9322
12.2757 | -34.5047 | -20.9408
12.2818 | -34.5052 | -20.9498
12.2879 | -34.5056 | -20.9586
12.294 | -34.5061 | -20.9675
12.3001 | -34.5066 | -20.9767
12.3062 | -34.5071 | -20.9857
12.3123 | -34.5076 | -20.9948
12.3184 | -34.5081 | -21.0036
12.3245 | -34.5086 | -21.0128
12.3306 | -34.509 | -21.0215
12.3367 | -34.5095 | -21.0303
12.3428 | -34.51 | -21.039
12.3489 | -34.5105 | -21.0475
12.355 | -34.5109 | -21.0564
12.3611 | -34.5114 | -21.0654
12.3672 | -34.5119 | -21.0742
12.3733 | -34.5124 | -21.0836
12.3794 | -34.5128 | -21.0919
12.3855 | -34.5133 | -21.1007
12.3916 | -34.5138 | -21.1098
12.3978 | -34.5142 | -21.1191
12.4039 | -34.5146 | -21.1281
12.41 | -34.5152 | -21.1372
12.4161 | -34.5157 | -21.1458
12.4222 | -34.5162 | -21.1545
12.4283 | -34.5167 | -21.1635
12.4344 | -34.5172 | -21.1724
12.4405 | -34.5176 | -21.1813
12.4466 | -34.5181 | -21.19
12.4527 | -34.5186 | -21.199
12.4588 | -34.5191 | -21.2079
12.4649 | -34.5195 | -21.2171
12.471 | -34.52 | -21.2266
12.4771 | -34.5205 | -21.2349
12.4832 | -34.521 | -21.2438
12.4893 | -34.5214 | -21.2523
12.4954 | -34.522 | -21.2612
12.5015 | -34.5225 | -21.2698
12.5076 | -34.5229 | -21.2788
12.5137 | -34.5234 | -21.2881
12.5198 | -34.5238 | -21.2967
12.5259 | -34.5244 | -21.3053
12.532 | -34.5249 | -21.3143
12.5382 | -34.5253 | -21.323
12.5443 | -34.5258 | -21.3319
12.5504 | -34.5262 | -21.3409
12.5565 | -34.5268 | -21.3494
12.5626 | -34.5273 | -21.3586
12.5687 | -34.5277 | -21.3676
12.5748 | -34.5282 | -21.3764
12.5809 | -34.5287 | -21.3849
12.587 | -34.5291 | -21.3941
12.5931 | -34.5296 | -21.4029
12.5992 | -34.5301 | -21.4119
12.6053 | -34.5306 | -21.4207
12.6114 | -34.5311 | -21.4296
12.6175 | -34.5316 | -21.4386
12.6236 | -34.532 | -21.4472
12.6297 | -34.5326 | -21.4559
12.6358 | -34.5331 | -21.4644
12.6419 | -34.5336 | -21.4735
12.648 | -34.5341 | -21.4827
12.6541 | -34.5345 | -21.4916
12.6602 | -34.535 | -21.5
12.6663 | -34.5355 | -21.5088
12.6724 | -34.536 | -21.5177
12.6785 | -34.5364 | -21.5269
12.6847 | -34.5369 | -21.5357
12.6908 | -34.5374 | -21.5444
12.6969 | -34.5379 | -21.5529
12.703 | -34.5384 | -21.5618
12.7091 | -34.5389 | -21.5708
12.7152 | -34.5393 | -21.5796
12.7213 | -34.5398 | -21.5886
12.7274 | -34.5403 | -21.5974
12.7335 | -34.5408 | -21.6056
12.7396 | -34.5412 | -21.614
12.7457 | -34.5417 | -21.6229
12.7518 | -34.5422 | -21.6316
12.7579 | -34.5427 | -21.6409
12.764 | -34.5431 | -21.6497
12.7701 | -34.5435 | -21.6581
12.7762 | -34.544 | -21.6668
12.7823 | -34.5445 | -21.6756
12.7884 | -34.5449 | -21.6842
12.7945 | -34.5454 | -21.6925
12.8006 | -34.5459 | -21.7007
12.8067 | -34.5463 | -21.7098
12.8128 | -34.5467 | -21.7181
12.8189 | -34.5472 | -21.7267
12.8251 | -34.5476 | -21.7353
12.8312 | -34.548 | -21.744
12.8373 | -34.5485 | -21.7526
12.8434 | -34.549 | -21.7612
12.8495 | -34.5494 | -21.7699
12.8556 | -34.5499 | -21.7786
12.8617 | -34.5503 | -21.7871
12.8678 | -34.5508 | -21.7959
12.8739 | -34.5513 | -21.8042
12.88 | -34.5517 | -21.8131
12.8861 | -34.5522 | -21.8218
12.8922 | -34.5526 | -21.8305
12.8983 | -34.5532 | -21.8389
12.9044 | -34.5537 | -21.8475
12.9105 | -34.5542 | -21.856
12.9166 | -34.5547 | -21.865
12.9227 | -34.5551 | -21.8742
12.9288 | -34.5557 | -21.8828
12.9349 | -34.5561 | -21.8917
12.941 | -34.5567 | -21.9006
12.9471 | -34.5572 | -21.9093
12.9532 | -34.5577 | -21.9179
12.9593 | -34.5582 | -21.9269
12.9654 | -34.5586 | -21.9357
12.9716 | -34.5591 | -21.9445
12.9777 | -34.5595 | -21.9527
12.9838 | -34.56 | -21.9612
12.9899 | -34.5604 | -21.9699
12.996 | -34.5608 | -21.9788
13.0021 | -34.5613 | -21.9876
13.0082 | -34.5618 | -21.9963
13.0143 | -34.5622 | -22.0045
13.0204 | -34.5626 | -22.0129
13.0265 | -34.5631 | -22.0213
13.0326 | -34.5635 | -22.0301
13.0387 | -34.564 | -22.039
13.0448 | -34.5645 | -22.0475
13.0509 | -34.565 | -22.056
13.057 | -34.5655 | -22.0643
13.0631 | -34.566 | -22.0728
13.0692 | -34.5665 | -22.0815
13.0753 | -34.567 | -22.0901
13.0814 | -34.5674 | -22.0985
13.0875 | -34.5679 | -22.1066
13.0936 | -34.5684 | -22.115
13.0997 | -34.5688 | -22.1231
13.1058 | -34.5693 | -22.1311
13.112 | -34.5697 | -22.1397
13.1181 | -34.5702 | -22.1481
13.1242 | -34.5706 | -22.1567
13.1303 | -34.5711 | -22.1655
13.1364 | -34.5716 | -22.1742
13.1425 | -34.5721 | -22.1827
13.1486 | -34.5725 | -22.191
13.1547 | -34.573 | -22.1992
13.1608 | -34.5735 | -22.2075
13.1669 | -34.574 | -22.216
13.173 | -34.5744 | -22.2247
13.1791 | -34.5749 | -22.2332
13.1852 | -34.5753 | -22.242
13.1913 | -34.5757 | -22.2506
13.1974 | -34.5762 | -22.2592
13.2035 | -34.5766 | -22.2676
13.2096 | -34.5772 | -22.2761
13.2157 | -34.5776 | -22.2845
13.2218 | -34.5781 | -22.2931
13.2279 | -34.5786 | -22.3015
13.234 | -34.5791 | -22.3102
13.2401 | -34.5795 | -22.3188
13.2462 | -34.58 | -22.3273
13.2524 | -34.5804 | -22.3355
13.2585 | -34.5809 | -22.3438
13.2646 | -34.5814 | -22.3524
13.2707 | -34.5818 | -22.3611
13.2768 | -34.5823 | -22.3701
13.2829 | -34.5828 | -22.3788
13.289 | -34.5832 | -22.3875
13.2951 | -34.5837 | -22.3956
13.3012 | -34.5842 | -22.404
13.3073 | -34.5846 | -22.413
13.3134 | -34.585 | -22.4214
13.3195 | -34.5855 | -22.4298
13.3256 | -34.586 | -22.4385
13.3317 | -34.5865 | -22.4467
13.3378 | -34.587 | -22.4554
13.3439 | -34.5874 | -22.4639
13.35 | -34.5879 | -22.4725
13.3561 | -34.5883 | -22.4811
13.3622 | -34.5887 | -22.4897
13.3683 | -34.5892 | -22.498
13.3744 | -34.5897 | -22.5063
13.3805 | -34.5901 | -22.5148
13.3866 | -34.5906 | -22.5228
13.3927 | -34.591 | -22.5314
13.3989 | -34.5915 | -22.5395
13.405 | -34.592 | -22.548
13.4111 | -34.5925 | -22.5565
13.4172 | -34.593 | -22.5648
13.4233 | -34.5934 | -22.5733
13.4294 | -34.5939 | -22.5814
13.4355 | -34.5944 | -22.5904
13.4416 | -34.5948 | -22.599
13.4477 | -34.5952 | -22.6075
13.4538 | -34.5956 | -22.6158
13.4599 | -34.5961 | -22.6245
13.466 | -34.5966 | -22.6335
13.4721 | -34.5971 | -22.6417
13.4782 | -34.5976 | -22.6501
13.4843 | -34.5981 | -22.6587
13.4904 | -34.5986 | -22.6665
13.4965 | -34.599 | -22.6753
13.5026 | -34.5995 | -22.6837
13.5087 | -34.6 | -22.6918
13.5148 | -34.6004 | -22.7003
13.5209 | -34.6009 | -22.7085
13.527 | -34.6013 | -22.7175
13.5331 | -34.6018 | -22.7259
13.5393 | -34.6023 | -22.7345
13.5454 | -34.6027 | -22.7429
13.5515 | -34.6032 | -22.7513
13.5576 | -34.6037 | -22.7599
13.5637 | -34.6042 | -22.7682
13.5698 | -34.6047 | -22.7768
13.5759 | -34.6051 | -22.7848
13.582 | -34.6056 | -22.7933
13.5881 | -34.606 | -22.8018
13.5942 | -34.6065 | -22.8102
13.6003 | -34.607 | -22.8187
13.6064 | -34.6074 | -22.8273
13.6125 | -34.6079 | -22.8362
13.6186 | -34.6084 | -22.8445
13.6247 | -34.6089 | -22.8529
13.6308 | -34.6093 | -22.861
13.6369 | -34.6098 | -22.8693
13.643 | -34.6103 | -22.8775
13.6491 | -34.6108 | -22.886
13.6552 | -34.6112 | -22.8942
13.6613 | -34.6117 | -22.903
13.6674 | -34.6122 | -22.9113
13.6735 | -34.6126 | -22.9199
13.6796 | -34.6131 | -22.9281
13.6858 | -34.6135 | -22.9364
13.6919 | -34.614 | -22.9448
13.698 | -34.6144 | -22.9533
13.7041 | -34.6148 | -22.9621
13.7102 | -34.6153 | -22.9706
13.7163 | -34.6159 | -22.9792
13.7224 | -34.6163 | -22.9873
13.7285 | -34.6168 | -22.9958
13.7346 | -34.6172 | -23.0041
13.7407 | -34.6177 | -23.0127
13.7468 | -34.6181 | -23.0213
13.7529 | -34.6186 | -23.0297
13.759 | -34.619 | -23.0382
13.7651 | -34.6195 | -23.0466
13.7712 | -34.6199 | -23.0553
13.7773 | -34.6204 | -23.0638
13.7834 | -34.6209 | -23.0721
13.7895 | -34.6214 | -23.0805
13.7956 | -34.6219 | -23.089
13.8017 | -34.6224 | -23.0971
13.8078 | -34.6229 | -23.1056
13.8139 | -34.6234 | -23.1139
13.82 | -34.6238 | -23.122
13.8262 | -34.6244 | -23.1304
13.8323 | -34.6249 | -23.1392
13.8384 | -34.6253 | -23.1477
13.8445 | -34.6258 | -23.1565
13.8506 | -34.6263 | -23.1652
13.8567 | -34.6268 | -23.1739
13.8628 | -34.6272 | -23.1824
13.8689 | -34.6276 | -23.1908
13.875 | -34.6282 | -23.199
13.8811 | -34.6287 | -23.2074
13.8872 | -34.6291 | -23.2159
13.8933 | -34.6297 | -23.2243
13.8994 | -34.6301 | -23.2324
13.9055 | -34.6306 | -23.2409
13.9116 | -34.6311 | -23.249
13.9177 | -34.6316 | -23.2571
13.9238 | -34.6321 | -23.2657
13.9299 | -34.6326 | -23.2743
13.936 | -34.633 | -23.2825
13.9421 | -34.6335 | -23.2909
13.9482 | -34.634 | -23.2991
13.9543 | -34.6345 | -23.3075
13.9604 | -34.635 | -23.3161
13.9665 | -34.6355 | -23.3247
13.9727 | -34.6359 | -23.3329
13.9788 | -34.6364 | -23.3416
13.9849 | -34.6368 | -23.35
13.991 | -34.6373 | -23.3587
13.9971 | -34.6377 | -23.3674
14.0032 | -34.6382 | -23.3761
14.0093 | -34.6386 | -23.3845
14.0154 | -34.6391 | -23.3931
14.0215 | -34.6395 | -23.4012
14.0276 | -34.64 | -23.4095
14.0337 | -34.6406 | -23.4181
14.0398 | -34.641 | -23.4263
14.0459 | -34.6415 | -23.435
14.052 | -34.6419 | -23.4435
14.0581 | -34.6424 | -23.4516
14.0642 | -34.6429 | -23.46
14.0703 | -34.6433 | -23.4685
14.0764 | -34.6438 | -23.4773
14.0825 | -34.6443 | -23.4857
14.0886 | -34.6447 | -23.4937
14.0947 | -34.6452 | -23.5017
14.1008 | -34.6457 | -23.5104
14.1069 | -34.6462 | -23.5187
14.1131 | -34.6466 | -23.5273
14.1192 | -34.6471 | -23.5354
14.1253 | -34.6475 | -23.5435
14.1314 | -34.648 | -23.5519
14.1375 | -34.6486 | -23.5604
14.1436 | -34.6491 | -23.5688
14.1497 | -34.6496 | -23.5771
14.1558 | -34.6501 | -23.5855
14.1619 | -34.6506 | -23.5939
14.168 | -34.6511 | -23.6022
14.1741 | -34.6516 | -23.6107
14.1802 | -34.6521 | -23.6189
14.1863 | -34.6526 | -23.6273
14.1924 | -34.653 | -23.6357
14.1985 | -34.6535 | -23.644
14.2046 | -34.654 | -23.6522
14.2107 | -34.6545 | -23.6602
14.2168 | -34.655 | -23.6692
14.2229 | -34.6555 | -23.6776
14.229 | -34.656 | -23.6861
14.2351 | -34.6565 | -23.6941
14.2412 | -34.6569 | -23.7026
14.2473 | -34.6575 | -23.711
14.2534 | -34.658 | -23.7193
14.2596 | -34.6584 | -23.7276
14.2657 | -34.6589 | -23.7362
14.2718 | -34.6594 | -23.7449
14.2779 | -34.6599 | -23.7531
14.284 | -34.6604 | -23.7615
14.2901 | -34.6609 | -23.7695
14.2962 | -34.6614 | -23.7777
14.3023 | -34.6619 | -23.7862
14.3084 | -34.6624 | -23.7947
14.3145 | -34.6629 | -23.8032
14.3206 | -34.6634 | -23.8116
14.3267 | -34.6638 | -23.8199
14.3328 | -34.6643 | -23.8283
14.3389 | -34.6648 | -23.8364
14.345 | -34.6652 | -23.8449
14.3511 | -34.6658 | -23.8531
14.3572 | -34.6662 | -23.8616
14.3633 | -34.6667 | -23.8699
14.3694 | -34.6672 | -23.8783
14.3755 | -34.6676 | -23.8863
14.3816 | -34.6681 | -23.8947
14.3877 | -34.6686 | -23.9029
14.3938 | -34.6691 | -23.9115
14.4 | -34.6696 | -23.9196
14.4061 | -34.67 | -23.9279
14.4122 | -34.6705 | -23.9364
14.4183 | -34.6709 | -23.9446
14.4244 | -34.6714 | -23.9529
14.4305 | -34.6719 | -23.9617
14.4366 | -34.6724 | -23.9702
14.4427 | -34.6729 | -23.9789
14.4488 | -34.6734 | -23.9871
14.4549 | -34.6739 | -23.9953
14.461 | -34.6744 | -24.0038
14.4671 | -34.6749 | -24.0122
14.4732 | -34.6754 | -24.0202
14.4793 | -34.6759 | -24.0293
14.4854 | -34.6764 | -24.0373
14.4915 | -34.6769 | -24.0453
14.4976 | -34.6775 | -24.0539
14.5037 | -34.678 | -24.0624
14.5098 | -34.6784 | -24.0706
14.5159 | -34.6789 | -24.0791
14.522 | -34.6794 | -24.0876
14.5281 | -34.6799 | -24.0958
14.5342 | -34.6804 | -24.1043
14.5403 | -34.6809 | -24.1127
14.5465 | -34.6814 | -24.1214
14.5526 | -34.6819 | -24.1301
14.5587 | -34.6824 | -24.1383
14.5648 | -34.6829 | -24.1465
14.5709 | -34.6834 | -24.155
14.577 | -34.6839 | -24.1634
14.5831 | -34.6843 | -24.1719
14.5892 | -34.6849 | -24.1799
14.5953 | -34.6854 | -24.1879
14.6014 | -34.6859 | -24.1963
14.6075 | -34.6864 | -24.2044
14.6136 | -34.6869 | -24.2125
14.6197 | -34.6874 | -24.221
14.6258 | -34.6879 | -24.2293
14.6319 | -34.6884 | -24.2374
14.638 | -34.6889 | -24.2458
14.6441 | -34.6894 | -24.2536
14.6502 | -34.6898 | -24.2617
14.6563 | -34.6903 | -24.27
14.6624 | -34.6909 | -24.2783
14.6685 | -34.6913 | -24.2862
14.6746 | -34.6919 | -24.2945
14.6807 | -34.6923 | -24.3027
14.6869 | -34.6928 | -24.311
14.693 | -34.6933 | -24.3192
14.6991 | -34.6938 | -24.3273
14.7052 | -34.6943 | -24.3356
14.7113 | -34.6948 | -24.3442
14.7174 | -34.6953 | -24.3524
14.7235 | -34.6958 | -24.3605
14.7296 | -34.6963 | -24.3687
14.7357 | -34.6968 | -24.3768
14.7418 | -34.6973 | -24.3848
14.7479 | -34.6978 | -24.3931
14.754 | -34.6984 | -24.4013
14.7601 | -34.6989 | -24.4096
14.7662 | -34.6994 | -24.418
14.7723 | -34.6999 | -24.4262
14.7784 | -34.7004 | -24.4346
14.7845 | -34.7009 | -24.443
14.7906 | -34.7014 | -24.4513
14.7967 | -34.7019 | -24.4594
14.8028 | -34.7024 | -24.4675
14.8089 | -34.7029 | -24.4757
14.815 | -34.7033 | -24.4838
14.8211 | -34.7038 | -24.4919
14.8272 | -34.7043 | -24.5003
14.8334 | -34.7048 | -24.5084
14.8395 | -34.7054 | -24.5166
14.8456 | -34.7059 | -24.5248
14.8517 | -34.7064 | -24.5331
14.8578 | -34.7069 | -24.5414
14.8639 | -34.7074 | -24.5497
14.87 | -34.7079 | -24.558
14.8761 | -34.7084 | -24.566
14.8822 | -34.7089 | -24.5742
14.8883 | -34.7094 | -24.5822
14.8944 | -34.7099 | -24.5904
14.9005 | -34.7104 | -24.5985
14.9066 | -34.711 | -24.6068
14.9127 | -34.7115 | -24.615
14.9188 | -34.712 | -24.6234
14.9249 | -34.7125 | -24.6317
14.931 | -34.713 | -24.6398
14.9371 | -34.7135 | -24.648
14.9432 | -34.714 | -24.6563
14.9493 | -34.7145 | -24.6645
14.9554 | -34.715 | -24.6726
14.9615 | -34.7155 | -24.6808
14.9676 | -34.716 | -24.6892
14.9738 | -34.7165 | -24.6973
14.9799 | -34.717 | -24.7058
14.986 | -34.7175 | -24.7136
14.9921 | -34.718 | -24.7217
14.9982 | -34.7185 | -24.7303
15.0043 | -34.719 | -24.7385
15.0104 | -34.7195 | -24.7467
15.0165 | -34.7201 | -24.7551
15.0226 | -34.7205 | -24.763
15.0287 | -34.721 | -24.7718
15.0348 | -34.7216 | -24.7802
15.0409 | -34.722 | -24.7887
15.047 | -34.7226 | -24.7972
15.0531 | -34.7232 | -24.8056
15.0592 | -34.7237 | -24.814
15.0653 | -34.7242 | -24.8221
15.0714 | -34.7247 | -24.83
15.0775 | -34.7253 | -24.8378
15.0836 | -34.7258 | -24.846
15.0897 | -34.7264 | -24.8542
15.0958 | -34.7269 | -24.8622
15.1019 | -34.7273 | -24.8704
15.108 | -34.7278 | -24.8787
15.1141 | -34.7284 | -24.8869
15.1203 | -34.7289 | -24.8952
15.1264 | -34.7295 | -24.9037
15.1325 | -34.73 | -24.9119
15.1386 | -34.7305 | -24.9201
15.1447 | -34.7311 | -24.9284
15.1508 | -34.7316 | -24.9368
15.1569 | -34.7321 | -24.9453
15.163 | -34.7327 | -24.953
15.1691 | -34.7332 | -24.9612
15.1752 | -34.7338 | -24.9694
15.1813 | -34.7343 | -24.9776
15.1874 | -34.7348 | -24.9859
15.1935 | -34.7354 | -24.9938
15.1996 | -34.7359 | -25.0022
15.2057 | -34.7364 | -25.0103
15.2118 | -34.7369 | -25.0184
15.2179 | -34.7374 | -25.0263
15.224 | -34.7379 | -25.0346
15.2301 | -34.7385 | -25.0425
15.2362 | -34.739 | -25.0507
15.2423 | -34.7395 | -25.059
15.2484 | -34.74 | -25.0671
15.2545 | -34.7406 | -25.0752
15.2607 | -34.7411 | -25.0834
15.2668 | -34.7417 | -25.0913
15.2729 | -34.7422 | -25.0994
15.279 | -34.7428 | -25.1078
15.2851 | -34.7433 | -25.116
15.2912 | -34.7438 | -25.1241
15.2973 | -34.7443 | -25.1321
15.3034 | -34.7449 | -25.1399
15.3095 | -34.7454 | -25.1478
15.3156 | -34.7459 | -25.1561
15.3217 | -34.7464 | -25.1639
15.3278 | -34.747 | -25.172
15.3339 | -34.7476 | -25.1801
15.34 | -34.7481 | -25.1879
15.3461 | -34.7486 | -25.1959
15.3522 | -34.7491 | -25.2039
15.3583 | -34.7497 | -25.2121
15.3644 | -34.7501 | -25.2205
15.3705 | -34.7507 | -25.2286
15.3766 | -34.7512 | -25.237
15.3827 | -34.7517 | -25.245
15.3888 | -34.7523 | -25.2529
15.3949 | -34.7528 | -25.2608
15.401 | -34.7533 | -25.2689
15.4072 | -34.7539 | -25.2768
15.4133 | -34.7544 | -25.2851
15.4194 | -34.7549 | -25.2934
15.4255 | -34.7555 | -25.3016
15.4316 | -34.756 | -25.3096
15.4377 | -34.7565 | -25.3176
15.4438 | -34.7571 | -25.3255
15.4499 | -34.7576 | -25.3335
15.456 | -34.758 | -25.3416
15.4621 | -34.7586 | -25.3498
15.4682 | -34.7592 | -25.3581
15.4743 | -34.7598 | -25.3661
15.4804 | -34.7603 | -25.3744
15.4865 | -34.7608 | -25.3825
15.4926 | -34.7613 | -25.3907
15.4987 | -34.7619 | -25.3989
15.5048 | -34.7624 | -25.4071
15.5109 | -34.7629 | -25.4149
15.517 | -34.7635 | -25.4232
15.5231 | -34.7641 | -25.4314
15.5292 | -34.7646 | -25.4392
15.5353 | -34.7651 | -25.4473
15.5414 | -34.7656 | -25.4556
15.5476 | -34.7662 | -25.4639
15.5537 | -34.7667 | -25.4718
15.5598 | -34.7673 | -25.4797
15.5659 | -34.7678 | -25.488
15.572 | -34.7684 | -25.4957
15.5781 | -34.7689 | -25.5038
15.5842 | -34.7695 | -25.5116
15.5903 | -34.77 | -25.5193
15.5964 | -34.7705 | -25.5272
15.6025 | -34.7711 | -25.5351
15.6086 | -34.7716 | -25.5434
15.6147 | -34.7721 | -25.5515
15.6208 | -34.7727 | -25.5595
15.6269 | -34.7733 | -25.5675
15.633 | -34.7738 | -25.5755
15.6391 | -34.7743 | -25.5837
15.6452 | -34.7749 | -25.5916
15.6513 | -34.7754 | -25.6001
15.6574 | -34.776 | -25.608
15.6635 | -34.7765 | -25.6161
15.6696 | -34.777 | -25.6241
15.6757 | -34.7776 | -25.6326
15.6818 | -34.7781 | -25.6411
15.688 | -34.7787 | -25.6492
15.6941 | -34.7792 | -25.6573
15.7002 | -34.7797 | -25.6652
15.7063 | -34.7802 | -25.6735
15.7124 | -34.7808 | -25.6817
15.7185 | -34.7813 | -25.6895
15.7246 | -34.7818 | -25.6973
15.7307 | -34.7824 | -25.7056
15.7368 | -34.7829 | -25.714
15.7429 | -34.7834 | -25.722
15.749 | -34.7839 | -25.73
15.7551 | -34.7844 | -25.7379
15.7612 | -34.7849 | -25.7455
15.7673 | -34.7854 | -25.7536
15.7734 | -34.7859 | -25.7614
15.7795 | -34.7865 | -25.7698
15.7856 | -34.787 | -25.7775
15.7917 | -34.7876 | -25.7854
15.7978 | -34.7881 | -25.7934
15.8039 | -34.7886 | -25.8015
15.81 | -34.7891 | -25.8093
15.8161 | -34.7897 | -25.8172
15.8222 | -34.7902 | -25.8253
15.8283 | -34.7908 | -25.8332
15.8345 | -34.7913 | -25.8412
15.8406 | -34.7919 | -25.8491
15.8467 | -34.7925 | -25.8572
15.8528 | -34.793 | -25.8651
15.8589 | -34.7936 | -25.8729
15.865 | -34.7941 | -25.8808
15.8711 | -34.7947 | -25.8889
15.8772 | -34.7953 | -25.8972
15.8833 | -34.7958 | -25.9058
15.8894 | -34.7964 | -25.9134
15.8955 | -34.7969 | -25.9215
15.9016 | -34.7975 | -25.9295
15.9077 | -34.798 | -25.9374
15.9138 | -34.7986 | -25.9452
15.9199 | -34.7991 | -25.9534
15.926 | -34.7997 | -25.9617
15.9321 | -34.8003 | -25.9697
15.9382 | -34.8008 | -25.9775
15.9443 | -34.8014 | -25.9858
15.9504 | -34.8019 | -25.9937
15.9565 | -34.8025 | -26.0017
15.9626 | -34.8031 | -26.0098
15.9687 | -34.8037 | -26.0177
15.9749 | -34.8043 | -26.0259
15.981 | -34.8048 | -26.0338
15.9871 | -34.8053 | -26.0419
15.9932 | -34.8058 | -26.05
15.9993 | -34.8064 | -26.058
16.0054 | -34.8069 | -26.0656
16.0115 | -34.8075 | -26.0735
16.0176 | -34.8081 | -26.0817
16.0237 | -34.8086 | -26.0899
16.0298 | -34.8092 | -26.0972
16.0359 | -34.8097 | -26.1051
16.042 | -34.8103 | -26.1131
16.0481 | -34.8108 | -26.1209
16.0542 | -34.8114 | -26.1291
16.0603 | -34.812 | -26.1371
16.0664 | -34.8125 | -26.1452
16.0725 | -34.8131 | -26.1533
16.0786 | -34.8137 | -26.1615
16.0847 | -34.8142 | -26.1692
16.0908 | -34.8147 | -26.1772
16.0969 | -34.8153 | -26.185
16.103 | -34.8158 | -26.1931
16.1091 | -34.8164 | -26.2008
16.1152 | -34.8169 | -26.2086
16.1214 | -34.8175 | -26.2163
16.1275 | -34.8181 | -26.2243
16.1336 | -34.8187 | -26.2324
16.1397 | -34.8193 | -26.2403
16.1458 | -34.8199 | -26.2479
16.1519 | -34.8204 | -26.2559
16.158 | -34.821 | -26.2641
16.1641 | -34.8216 | -26.2722
16.1702 | -34.8221 | -26.2803
16.1763 | -34.8227 | -26.2882
16.1824 | -34.8233 | -26.296
16.1885 | -34.8239 | -26.3041
16.1946 | -34.8245 | -26.3122
16.2007 | -34.825 | -26.3206
16.2068 | -34.8255 | -26.3283
16.2129 | -34.826 | -26.3364
16.219 | -34.8265 | -26.3442
16.2251 | -34.8271 | -26.3522
16.2312 | -34.8276 | -26.3602
16.2373 | -34.8281 | -26.3681
16.2434 | -34.8287 | -26.3759
16.2495 | -34.8293 | -26.3838
16.2556 | -34.8299 | -26.3918
16.2618 | -34.8304 | -26.3996
16.2679 | -34.831 | -26.4077
16.274 | -34.8316 | -26.4158
16.2801 | -34.8321 | -26.4234
16.2862 | -34.8327 | -26.4313
16.2923 | -34.8333 | -26.4395
16.2984 | -34.8338 | -26.4474
16.3045 | -34.8344 | -26.4552
16.3106 | -34.8349 | -26.4629
16.3167 | -34.8354 | -26.471
16.3228 | -34.836 | -26.479
16.3289 | -34.8365 | -26.4871
16.335 | -34.8371 | -26.4952
16.3411 | -34.8376 | -26.5029
16.3472 | -34.8382 | -26.511
16.3533 | -34.8387 | -26.5192
16.3594 | -34.8392 | -26.5268
16.3655 | -34.8398 | -26.5351
16.3716 | -34.8403 | -26.5429
16.3777 | -34.8409 | -26.5509
16.3838 | -34.8414 | -26.5586
16.3899 | -34.8419 | -26.5663
16.396 | -34.8425 | -26.5744
16.4021 | -34.843 | -26.5824
16.4083 | -34.8436 | -26.5902
16.4144 | -34.8442 | -26.598
16.4205 | -34.8448 | -26.6062
16.4266 | -34.8453 | -26.614
16.4327 | -34.8459 | -26.6218
16.4388 | -34.8464 | -26.6298
16.4449 | -34.847 | -26.6378
16.451 | -34.8475 | -26.6458
16.4571 | -34.848 | -26.654
16.4632 | -34.8485 | -26.6617
16.4693 | -34.8491 | -26.6698
16.4754 | -34.8496 | -26.6779
16.4815 | -34.8502 | -26.6859
16.4876 | -34.8508 | -26.6937
16.4937 | -34.8513 | -26.7016
16.4998 | -34.8519 | -26.7096
16.5059 | -34.8525 | -26.7174
16.512 | -34.853 | -26.7258
16.5181 | -34.8536 | -26.7339
16.5242 | -34.8542 | -26.7418
16.5303 | -34.8547 | -26.7498
16.5364 | -34.8552 | -26.7575
16.5425 | -34.8558 | -26.7655
16.5487 | -34.8563 | -26.7734
16.5548 | -34.8568 | -26.7815
16.5609 | -34.8574 | -26.7897
16.567 | -34.858 | -26.7977
16.5731 | -34.8585 | -26.8058
16.5792 | -34.8591 | -26.8138
16.5853 | -34.8597 | -26.8215
16.5914 | -34.8603 | -26.8295
16.5975 | -34.8608 | -26.8376
16.6036 | -34.8614 | -26.8456
16.6097 | -34.862 | -26.8534
16.6158 | -34.8625 | -26.8613
16.6219 | -34.8631 | -26.869
16.628 | -34.8636 | -26.877
16.6341 | -34.8642 | -26.8848
16.6402 | -34.8648 | -26.8925
16.6463 | -34.8653 | -26.9004
16.6524 | -34.8659 | -26.9083
16.6585 | -34.8665 | -26.9163
16.6646 | -34.867 | -26.9244
16.6707 | -34.8676 | -26.9324
16.6768 | -34.8682 | -26.9401
16.6829 | -34.8688 | -26.9481
16.689 | -34.8693 | -26.9561
16.6952 | -34.8698 | -26.9642
16.7013 | -34.8704 | -26.9722
16.7074 | -34.871 | -26.9801
16.7135 | -34.8715 | -26.9881
16.7196 | -34.8721 | -26.9959
16.7257 | -34.8727 | -27.0037
16.7318 | -34.8732 | -27.0117
16.7379 | -34.8738 | -27.0196
16.744 | -34.8743 | -27.0272
16.7501 | -34.8749 | -27.0353
16.7562 | -34.8755 | -27.0431
16.7623 | -34.876 | -27.0511
16.7684 | -34.8766 | -27.0592
16.7745 | -34.8772 | -27.0668
16.7806 | -34.8777 | -27.0749
16.7867 | -34.8783 | -27.0828
16.7928 | -34.8788 | -27.0908
16.7989 | -34.8793 | -27.0987
16.805 | -34.8799 | -27.1066
16.8111 | -34.8804 | -27.1145
16.8172 | -34.881 | -27.1228
16.8233 | -34.8816 | -27.1309
16.8294 | -34.8821 | -27.1388
16.8356 | -34.8827 | -27.147
16.8417 | -34.8832 | -27.1548
16.8478 | -34.8838 | -27.1625
16.8539 | -34.8844 | -27.1704
16.86 | -34.885 | -27.1783
16.8661 | -34.8855 | -27.1862
16.8722 | -34.8861 | -27.1942
16.8783 | -34.8866 | -27.2021
16.8844 | -34.8872 | -27.21
16.8905 | -34.8877 | -27.2179
16.8966 | -34.8883 | -27.2259
16.9027 | -34.8889 | -27.2337
16.9088 | -34.8894 | -27.2421
16.9149 | -34.89 | -27.2503
16.921 | -34.8906 | -27.258
16.9271 | -34.8911 | -27.2663
16.9332 | -34.8917 | -27.2739
16.9393 | -34.8923 | -27.2822
16.9454 | -34.8928 | -27.2899
16.9515 | -34.8933 | -27.298
16.9576 | -34.8939 | -27.3057
16.9637 | -34.8944 | -27.3137
16.9698 | -34.895 | -27.3212
16.9759 | -34.8956 | -27.3289
16.9821 | -34.8961 | -27.337
16.9882 | -34.8967 | -27.345
16.9943 | -34.8972 | -27.353
17.0004 | -34.8978 | -27.3609
17.0065 | -34.8983 | -27.3687
17.0126 | -34.8989 | -27.3766
17.0187 | -34.8994 | -27.3846
17.0248 | -34.9 | -27.3923
17.0309 | -34.9006 | -27.4003
17.037 | -34.9011 | -27.4082
17.0431 | -34.9017 | -27.4158
17.0492 | -34.9023 | -27.4237
17.0553 | -34.9028 | -27.4316
17.0614 | -34.9033 | -27.4396
17.0675 | -34.9039 | -27.4478
17.0736 | -34.9044 | -27.4556
17.0797 | -34.9049 | -27.463
17.0858 | -34.9055 | -27.4711
17.0919 | -34.906 | -27.479
17.098 | -34.9066 | -27.4868
17.1041 | -34.9072 | -27.4943
17.1102 | -34.9077 | -27.5024
17.1163 | -34.9083 | -27.5105
17.1225 | -34.9088 | -27.5182
17.1286 | -34.9094 | -27.5259
17.1347 | -34.91 | -27.5337
17.1408 | -34.9105 | -27.5416
17.1469 | -34.9111 | -27.5494
17.153 | -34.9117 | -27.5571
17.1591 | -34.9122 | -27.5646
17.1652 | -34.9127 | -27.5725
17.1713 | -34.9133 | -27.5802
17.1774 | -34.9138 | -27.5881
17.1835 | -34.9144 | -27.5963
17.1896 | -34.9149 | -27.6039
17.1957 | -34.9154 | -27.6118
17.2018 | -34.916 | -27.6198
17.2079 | -34.9165 | -27.6275
17.214 | -34.9171 | -27.6353
17.2201 | -34.9176 | -27.6433
17.2262 | -34.9181 | -27.651
17.2323 | -34.9187 | -27.659
17.2384 | -34.9193 | -27.6668
17.2445 | -34.9198 | -27.6744
17.2506 | -34.9203 | -27.6825
17.2567 | -34.9209 | -27.6903
17.2628 | -34.9214 | -27.6984
17.269 | -34.922 | -27.7061
17.2751 | -34.9225 | -27.7139
17.2812 | -34.9231 | -27.7218
17.2873 | -34.9237 | -27.7299
17.2934 | -34.9242 | -27.7379
17.2995 | -34.9247 | -27.7457
17.3056 | -34.9253 | -27.7534
17.3117 | -34.9258 | -27.7612
17.3178 | -34.9264 | -27.7689
17.3239 | -34.9269 | -27.7767
17.33 | -34.9275 | -27.7846
17.3361 | -34.928 | -27.7924
17.3422 | -34.9286 | -27.8001
17.3483 | -34.9291 | -27.8081
17.3544 | -34.9297 | -27.8157
17.3605 | -34.9302 | -27.8234
17.3666 | -34.9308 | -27.8314
17.3727 | -34.9313 | -27.839
17.3788 | -34.9319 | -27.8469
17.3849 | -34.9324 | -27.8546
17.391 | -34.933 | -27.8623
17.3971 | -34.9336 | -27.8702
17.4032 | -34.9341 | -27.8783
17.4094 | -34.9347 | -27.886
17.4155 | -34.9351 | -27.8937
17.4216 | -34.9357 | -27.9016
17.4277 | -34.9362 | -27.9096
17.4338 | -34.9368 | -27.9175
17.4399 | -34.9374 | -27.9253
17.446 | -34.9379 | -27.9329
17.4521 | -34.9385 | -27.9406
17.4582 | -34.939 | -27.9482
17.4643 | -34.9395 | -27.9561
17.4704 | -34.9401 | -27.9638
17.4765 | -34.9406 | -27.9716
17.4826 | -34.9412 | -27.9793
17.4887 | -34.9417 | -27.9876
17.4948 | -34.9423 | -27.9953
17.5009 | -34.9428 | -28.0029
17.507 | -34.9433 | -28.0109
17.5131 | -34.9439 | -28.0187
17.5192 | -34.9444 | -28.0264
17.5253 | -34.9449 | -28.0343
17.5314 | -34.9455 | -28.0422
17.5375 | -34.946 | -28.0502
17.5436 | -34.9465 | -28.0581
17.5497 | -34.9471 | -28.066
17.5559 | -34.9476 | -28.0736
17.562 | -34.9482 | -28.0813
17.5681 | -34.9487 | -28.0891
17.5742 | -34.9492 | -28.0969
17.5803 | -34.9498 | -28.1044
17.5864 | -34.9504 | -28.112
17.5925 | -34.951 | -28.1201
17.5986 | -34.9515 | -28.1276
17.6047 | -34.9521 | -28.1357
17.6108 | -34.9527 | -28.1432
17.6169 | -34.9532 | -28.1509
17.623 | -34.9538 | -28.1585
17.6291 | -34.9543 | -28.1661
17.6352 | -34.9549 | -28.1737
17.6413 | -34.9554 | -28.1817
17.6474 | -34.9559 | -28.1897
17.6535 | -34.9565 | -28.1976
17.6596 | -34.957 | -28.2052
17.6657 | -34.9575 | -28.2128
17.6718 | -34.9581 | -28.2208
17.6779 | -34.9586 | -28.2285
17.684 | -34.9592 | -28.2364
17.6901 | -34.9597 | -28.244
17.6963 | -34.9603 | -28.2515
17.7024 | -34.9608 | -28.2592
17.7085 | -34.9613 | -28.2666
17.7146 | -34.9619 | -28.2743
17.7207 | -34.9624 | -28.282
17.7268 | -34.963 | -28.2896
17.7329 | -34.9636 | -28.2976
17.739 | -34.9641 | -28.3055
17.7451 | -34.9647 | -28.3131
17.7512 | -34.9652 | -28.3211
17.7573 | -34.9657 | -28.3288
17.7634 | -34.9663 | -28.3363
17.7695 | -34.9668 | -28.3442
17.7756 | -34.9674 | -28.3519
17.7817 | -34.968 | -28.3595
17.7878 | -34.9685 | -28.3672
17.7939 | -34.9691 | -28.3749
17.8 | -34.9696 | -28.3824
17.8061 | -34.9701 | -28.3903
17.8122 | -34.9707 | -28.3981
17.8183 | -34.9712 | -28.406
17.8244 | -34.9717 | -28.4137
17.8305 | -34.9723 | -28.4215
17.8366 | -34.9729 | -28.4289
17.8428 | -34.9734 | -28.4369
17.8489 | -34.974 | -28.4445
17.855 | -34.9745 | -28.452
17.8611 | -34.9751 | -28.4595
17.8672 | -34.9756 | -28.4671
17.8733 | -34.9761 | -28.4749
17.8794 | -34.9767 | -28.4826
17.8855 | -34.9772 | -28.4901
17.8916 | -34.9778 | -28.4976
17.8977 | -34.9783 | -28.5053
17.9038 | -34.9789 | -28.5128
17.9099 | -34.9794 | -28.5205
17.916 | -34.9799 | -28.5282
17.9221 | -34.9805 | -28.5357
17.9282 | -34.981 | -28.5433
17.9343 | -34.9816 | -28.551
17.9404 | -34.9821 | -28.5584
17.9465 | -34.9826 | -28.5662
17.9526 | -34.9831 | -28.5741
17.9587 | -34.9837 | -28.5818
17.9648 | -34.9842 | -28.5893
17.9709 | -34.9848 | -28.5969
17.977 | -34.9853 | -28.6042
17.9832 | -34.9859 | -28.6119
17.9893 | -34.9864 | -28.6196
17.9954 | -34.9869 | -28.6274
18.0015 | -34.9875 | -28.6352
18.0076 | -34.988 | -28.6424
18.0137 | -34.9885 | -28.6502
18.0198 | -34.9891 | -28.6579
18.0259 | -34.9896 | -28.6656
18.032 | -34.9901 | -28.673
18.0381 | -34.9906 | -28.6805
18.0442 | -34.9912 | -28.6883
18.0503 | -34.9917 | -28.6961
18.0564 | -34.9922 | -28.7037
18.0625 | -34.9928 | -28.7113
18.0686 | -34.9933 | -28.7188
18.0747 | -34.9938 | -28.7264
18.0808 | -34.9944 | -28.7338
18.0869 | -34.9949 | -28.7415
18.093 | -34.9955 | -28.7492
18.0991 | -34.996 | -28.7571
18.1052 | -34.9966 | -28.765
18.1113 | -34.9971 | -28.7723
18.1174 | -34.9977 | -28.7798
18.1236 | -34.9982 | -28.7875
18.1297 | -34.9988 | -28.7949
18.1358 | -34.9993 | -28.8027
18.1419 | -34.9999 | -28.8106
18.148 | -35.0004 | -28.8183
18.1541 | -35.001 | -28.8257
18.1602 | -35.0016 | -28.8335
18.1663 | -35.0022 | -28.841
18.1724 | -35.0027 | -28.8489
18.1785 | -35.0032 | -28.8567
18.1846 | -35.0037 | -28.8642
18.1907 | -35.0043 | -28.8722
18.1968 | -35.0048 | -28.8798
18.2029 | -35.0053 | -28.887
18.209 | -35.0059 | -28.8951
18.2151 | -35.0065 | -28.9028
18.2212 | -35.007 | -28.9101
18.2273 | -35.0076 | -28.9176
18.2334 | -35.0082 | -28.9253
18.2395 | -35.0087 | -28.9328
18.2456 | -35.0092 | -28.9403
18.2517 | -35.0097 | -28.9481
18.2578 | -35.0102 | -28.9553
18.2639 | -35.0108 | -28.9632
18.2701 | -35.0114 | -28.9708
18.2762 | -35.012 | -28.9785
18.2823 | -35.0125 | -28.986
18.2884 | -35.0131 | -28.9936
18.2945 | -35.0136 | -29.0008
18.3006 | -35.0141 | -29.0084
18.3067 | -35.0147 | -29.0158
18.3128 | -35.0152 | -29.0234
18.3189 | -35.0158 | -29.0312
18.325 | -35.0163 | -29.0383
18.3311 | -35.0169 | -29.046
18.3372 | -35.0174 | -29.0531
18.3433 | -35.0179 | -29.0606
18.3494 | -35.0184 | -29.0682
18.3555 | -35.019 | -29.0756
18.3616 | -35.0196 | -29.0828
18.3677 | -35.0201 | -29.0902
18.3738 | -35.0206 | -29.0978
18.3799 | -35.0212 | -29.1052
18.386 | -35.0217 | -29.1125
18.3921 | -35.0222 | -29.12
18.3982 | -35.0228 | -29.127
18.4043 | -35.0233 | -29.1345
18.4105 | -35.0238 | -29.1422
18.4166 | -35.0243 | -29.1497
18.4227 | -35.0249 | -29.1574
18.4288 | -35.0254 | -29.165
18.4349 | -35.026 | -29.1724
18.441 | -35.0265 | -29.1799
18.4471 | -35.0271 | -29.1875
18.4532 | -35.0276 | -29.1952
18.4593 | -35.0282 | -29.2028
18.4654 | -35.0287 | -29.2099
18.4715 | -35.0292 | -29.2174
18.4776 | -35.0298 | -29.225
18.4837 | -35.0303 | -29.2326
18.4898 | -35.0309 | -29.2401
18.4959 | -35.0315 | -29.2478
18.502 | -35.032 | -29.2552
18.5081 | -35.0326 | -29.2628
18.5142 | -35.0331 | -29.2705
18.5203 | -35.0337 | -29.2777
18.5264 | -35.0342 | -29.285
18.5325 | -35.0347 | -29.2926
18.5386 | -35.0353 | -29.3002
18.5447 | -35.0359 | -29.3077
18.5508 | -35.0364 | -29.3152
18.557 | -35.0369 | -29.3228
18.5631 | -35.0374 | -29.3303
18.5692 | -35.038 | -29.3376
18.5753 | -35.0386 | -29.3449
18.5814 | -35.0391 | -29.3526
18.5875 | -35.0397 | -29.3603
18.5936 | -35.0402 | -29.3681
18.5997 | -35.0407 | -29.3758
18.6058 | -35.0412 | -29.3832
18.6119 | -35.0418 | -29.3908
18.618 | -35.0424 | -29.3985
18.6241 | -35.0429 | -29.4063
18.6302 | -35.0435 | -29.4141
18.6363 | -35.044 | -29.4213
18.6424 | -35.0445 | -29.4284
18.6485 | -35.045 | -29.4359
18.6546 | -35.0456 | -29.4431
18.6607 | -35.0461 | -29.4504
18.6668 | -35.0467 | -29.4577
18.6729 | -35.0472 | -29.4653
18.679 | -35.0478 | -29.4724
18.6851 | -35.0483 | -29.4799
18.6912 | -35.0488 | -29.4874
18.6974 | -35.0494 | -29.4949
18.7035 | -35.0499 | -29.5025
18.7096 | -35.0505 | -29.51
18.7157 | -35.051 | -29.5174
18.7218 | -35.0516 | -29.5247
18.7279 | -35.0521 | -29.532
18.734 | -35.0527 | -29.5395
18.7401 | -35.0533 | -29.5466
18.7462 | -35.0538 | -29.5534
18.7523 | -35.0544 | -29.5608
18.7584 | -35.055 | -29.5685
18.7645 | -35.0555 | -29.5759
18.7706 | -35.0561 | -29.5834
18.7767 | -35.0566 | -29.5908
18.7828 | -35.0571 | -29.5979
18.7889 | -35.0577 | -29.6052
18.795 | -35.0582 | -29.6126
18.8011 | -35.0588 | -29.6201
18.8072 | -35.0593 | -29.6275
18.8133 | -35.0599 | -29.6348
18.8194 | -35.0604 | -29.6419
18.8255 | -35.061 | -29.6492
18.8316 | -35.0615 | -29.6569
18.8377 | -35.062 | -29.6643
18.8439 | -35.0626 | -29.6717
18.85 | -35.0631 | -29.6792
18.8561 | -35.0636 | -29.6866
18.8622 | -35.0641 | -29.694
18.8683 | -35.0647 | -29.7011
18.8744 | -35.0652 | -29.7083
18.8805 | -35.0658 | -29.7157
18.8866 | -35.0664 | -29.7233
18.8927 | -35.0669 | -29.7303
18.8988 | -35.0675 | -29.7377
18.9049 | -35.068 | -29.7451
18.911 | -35.0686 | -29.7524
18.9171 | -35.0692 | -29.7597
18.9232 | -35.0697 | -29.7668
18.9293 | -35.0703 | -29.7743
18.9354 | -35.0708 | -29.7815
18.9415 | -35.0713 | -29.7889
18.9476 | -35.0718 | -29.7963
18.9537 | -35.0724 | -29.804
18.9598 | -35.073 | -29.8106
18.9659 | -35.0735 | -29.818
18.972 | -35.074 | -29.8252
18.9781 | -35.0746 | -29.8324
18.9843 | -35.0752 | -29.8397
18.9904 | -35.0757 | -29.8473
18.9965 | -35.0762 | -29.8545
19.0026 | -35.0768 | -29.8619
19.0087 | -35.0773 | -29.8693
19.0148 | -35.0779 | -29.8766
19.0209 | -35.0784 | -29.8843
19.027 | -35.079 | -29.8917
19.0331 | -35.0795 | -29.8991
19.0392 | -35.08 | -29.9066
19.0453 | -35.0806 | -29.9137
19.0514 | -35.0811 | -29.921
19.0575 | -35.0816 | -29.9282
19.0636 | -35.0822 | -29.935
19.0697 | -35.0827 | -29.9426
19.0758 | -35.0833 | -29.9498
19.0819 | -35.0838 | -29.9576
19.088 | -35.0844 | -29.9646
19.0941 | -35.085 | -29.9715
19.1002 | -35.0855 | -29.9784
19.1063 | -35.0861 | -29.9858
19.1124 | -35.0866 | -29.9931
19.1185 | -35.0872 | -30.0006
19.1246 | -35.0877 | -30.0077
19.1308 | -35.0883 | -30.0147
19.1369 | -35.0888 | -30.0216
19.143 | -35.0893 | -30.0291
19.1491 | -35.0899 | -30.0365
19.1552 | -35.0904 | -30.0441
19.1613 | -35.091 | -30.0513
19.1674 | -35.0915 | -30.0587
19.1735 | -35.092 | -30.0659
19.1796 | -35.0926 | -30.0731
19.1857 | -35.0931 | -30.0806
19.1918 | -35.0936 | -30.088
19.1979 | -35.0941 | -30.0954
19.204 | -35.0947 | -30.1024
19.2101 | -35.0952 | -30.1097
19.2162 | -35.0958 | -30.1171
19.2223 | -35.0963 | -30.1239
19.2284 | -35.0969 | -30.1314
19.2345 | -35.0974 | -30.1387
19.2406 | -35.098 | -30.146
19.2467 | -35.0984 | -30.1532
19.2528 | -35.099 | -30.1607
19.2589 | -35.0995 | -30.168
19.265 | -35.1 | -30.1752
19.2712 | -35.1005 | -30.1824
19.2773 | -35.1011 | -30.1894
19.2834 | -35.1016 | -30.1965
19.2895 | -35.1022 | -30.2039
19.2956 | -35.1027 | -30.2113
19.3017 | -35.1032 | -30.2185
19.3078 | -35.1038 | -30.2257
19.3139 | -35.1044 | -30.2326
19.32 | -35.1049 | -30.2398
19.3261 | -35.1054 | -30.2471
19.3322 | -35.106 | -30.2542
19.3383 | -35.1066 | -30.2611
19.3444 | -35.1072 | -30.2681
19.3505 | -35.1077 | -30.2753
19.3566 | -35.1082 | -30.2825
19.3627 | -35.1088 | -30.2899
19.3688 | -35.1093 | -30.2973
19.3749 | -35.1098 | -30.3047
19.381 | -35.1104 | -30.3113
19.3871 | -35.1109 | -30.3186
19.3932 | -35.1114 | -30.3258
19.3993 | -35.1119 | -30.333
19.4054 | -35.1125 | -30.3406
19.4115 | -35.1131 | -30.3478
19.4177 | -35.1137 | -30.355
19.4238 | -35.1142 | -30.3621
19.4299 | -35.1147 | -30.3691
19.436 | -35.1153 | -30.3765
19.4421 | -35.1159 | -30.3835
19.4482 | -35.1165 | -30.3907
19.4543 | -35.1171 | -30.3972
19.4604 | -35.1176 | -30.4043
19.4665 | -35.1181 | -30.4114
19.4726 | -35.1187 | -30.4184
19.4787 | -35.1192 | -30.4255
19.4848 | -35.1197 | -30.4326
19.4909 | -35.1203 | -30.4397
19.497 | -35.1208 | -30.4469
19.5031 | -35.1214 | -30.4543
19.5092 | -35.1219 | -30.4614
19.5153 | -35.1225 | -30.4688
19.5214 | -35.123 | -30.4755
19.5275 | -35.1236 | -30.4825
19.5336 | -35.1241 | -30.4895
19.5397 | -35.1247 | -30.4966
19.5458 | -35.1253 | -30.5038
19.5519 | -35.1258 | -30.5109
19.5581 | -35.1264 | -30.5177
19.5642 | -35.1269 | -30.5248
19.5703 | -35.1274 | -30.5317
19.5764 | -35.128 | -30.5389
19.5825 | -35.1285 | -30.5461
19.5886 | -35.129 | -30.5531
19.5947 | -35.1296 | -30.56
19.6008 | -35.1301 | -30.5672
19.6069 | -35.1306 | -30.5744
19.613 | -35.1312 | -30.5818
19.6191 | -35.1317 | -30.5888
19.6252 | -35.1323 | -30.5959
19.6313 | -35.1328 | -30.6026
19.6374 | -35.1333 | -30.61
19.6435 | -35.1338 | -30.6171
19.6496 | -35.1343 | -30.6243
19.6557 | -35.1348 | -30.631
19.6618 | -35.1354 | -30.6379
19.6679 | -35.1359 | -30.6451
19.674 | -35.1365 | -30.652
19.6801 | -35.137 | -30.659
19.6862 | -35.1375 | -30.6661
19.6923 | -35.1381 | -30.6729
19.6984 | -35.1386 | -30.6795
19.7046 | -35.1392 | -30.6867
19.7107 | -35.1398 | -30.6939
19.7168 | -35.1403 | -30.7015
19.7229 | -35.1408 | -30.7089
19.729 | -35.1413 | -30.7161
19.7351 | -35.1418 | -30.7228
19.7412 | -35.1424 | -30.7298
19.7473 | -35.1429 | -30.737
19.7534 | -35.1435 | -30.7443
19.7595 | -35.144 | -30.7515
19.7656 | -35.1446 | -30.7587
19.7717 | -35.1451 | -30.7654
19.7778 | -35.1456 | -30.7726
19.7839 | -35.1461 | -30.7793
19.79 | -35.1467 | -30.7864
19.7961 | -35.1472 | -30.7937
19.8022 | -35.1477 | -30.8008
19.8083 | -35.1483 | -30.8079
19.8144 | -35.1489 | -30.815
19.8205 | -35.1494 | -30.8222
19.8266 | -35.1499 | -30.8289
19.8327 | -35.1505 | -30.836
19.8388 | -35.151 | -30.8432
19.845 | -35.1516 | -30.85
19.8511 | -35.1521 | -30.8574
19.8572 | -35.1527 | -30.8647
19.8633 | -35.1532 | -30.8714
19.8694 | -35.1538 | -30.8787
19.8755 | -35.1544 | -30.8853
19.8816 | -35.155 | -30.8922
19.8877 | -35.1555 | -30.8992
19.8938 | -35.156 | -30.9064
19.8999 | -35.1565 | -30.9133
19.906 | -35.1571 | -30.9202
19.9121 | -35.1576 | -30.9268
19.9182 | -35.1581 | -30.9338
19.9243 | -35.1587 | -30.9413
19.9304 | -35.1592 | -30.9487
19.9365 | -35.1597 | -30.9561
19.9426 | -35.1603 | -30.9628
19.9487 | -35.1609 | -30.9697
19.9548 | -35.1614 | -30.9772
19.9609 | -35.1619 | -30.9842
19.967 | -35.1624 | -30.9911
19.9731 | -35.163 | -30.9981
19.9792 | -35.1635 | -31.0048
19.9853 | -35.1641 | -31.0118
19.9915 | -35.1647 | -31.0189
19.9976 | -35.1652 | -31.0258
20.0037 | -35.1658 | -31.0331
20.0098 | -35.1663 | -31.0401
20.0159 | -35.1669 | -31.0467
20.022 | -35.1674 | -31.054
20.0281 | -35.168 | -31.0616
20.0342 | -35.1685 | -31.0683
20.0403 | -35.169 | -31.075
20.0464 | -35.1695 | -31.0825
20.0525 | -35.1701 | -31.0893
20.0586 | -35.1706 | -31.0962
20.0647 | -35.1712 | -31.1033
20.0708 | -35.1718 | -31.1108
20.0769 | -35.1723 | -31.1176
20.083 | -35.1728 | -31.1241
20.0891 | -35.1733 | -31.1312
20.0952 | -35.1739 | -31.1381
20.1013 | -35.1744 | -31.1449
20.1074 | -35.1749 | -31.1519
20.1135 | -35.1755 | -31.1589
20.1196 | -35.176 | -31.1657
20.1257 | -35.1765 | -31.1727
20.1319 | -35.177 | -31.1796
20.138 | -35.1776 | -31.1863
20.1441 | -35.1781 | -31.1932
20.1502 | -35.1786 | -31.2006
20.1563 | -35.1792 | -31.207
20.1624 | -35.1797 | -31.2138
20.1685 | -35.1803 | -31.2206
20.1746 | -35.1808 | -31.2274
20.1807 | -35.1813 | -31.2342
20.1868 | -35.1818 | -31.2412
20.1929 | -35.1824 | -31.248
20.199 | -35.1829 | -31.2553
20.2051 | -35.1834 | -31.2621
20.2112 | -35.184 | -31.2689
20.2173 | -35.1845 | -31.2759
20.2234 | -35.185 | -31.2826
20.2295 | -35.1856 | -31.2894
20.2356 | -35.1861 | -31.2964
20.2417 | -35.1867 | -31.3029
20.2478 | -35.1872 | -31.31
20.2539 | -35.1877 | -31.3171
20.26 | -35.1881 | -31.3239
20.2661 | -35.1887 | -31.3312
20.2723 | -35.1892 | -31.3381
20.2784 | -35.1897 | -31.3451
20.2845 | -35.1902 | -31.3519
20.2906 | -35.1908 | -31.359
20.2967 | -35.1914 | -31.3656
20.3028 | -35.1919 | -31.3724
20.3089 | -35.1924 | -31.3796
20.315 | -35.193 | -31.3865
20.3211 | -35.1935 | -31.3934
20.3272 | -35.194 | -31.4003
20.3333 | -35.1946 | -31.4066
20.3394 | -35.1951 | -31.4131
20.3455 | -35.1957 | -31.4201
20.3516 | -35.1962 | -31.427
20.3577 | -35.1968 | -31.4337
20.3638 | -35.1974 | -31.4406
20.3699 | -35.198 | -31.4468
20.376 | -35.1986 | -31.4542
20.3821 | -35.1991 | -31.461
20.3882 | -35.1996 | -31.468
20.3943 | -35.2002 | -31.4746
20.4004 | -35.2007 | -31.4813
20.4065 | -35.2013 | -31.4882
20.4126 | -35.2018 | -31.4948
20.4188 | -35.2024 | -31.5019
20.4249 | -35.203 | -31.509
20.431 | -35.2035 | -31.5157
20.4371 | -35.2041 | -31.5218
20.4432 | -35.2046 | -31.5287
20.4493 | -35.2051 | -31.5355
20.4554 | -35.2056 | -31.5422
20.4615 | -35.2062 | -31.549
20.4676 | -35.2068 | -31.5556
20.4737 | -35.2073 | -31.5623
20.4798 | -35.2079 | -31.5691
20.4859 | -35.2084 | -31.5758
20.492 | -35.209 | -31.5826
20.4981 | -35.2095 | -31.5892
20.5042 | -35.21 | -31.596
20.5103 | -35.2106 | -31.6025
20.5164 | -35.2111 | -31.6089
20.5225 | -35.2116 | -31.6164
20.5286 | -35.2122 | -31.624
20.5347 | -35.2127 | -31.6306
20.5408 | -35.2133 | -31.6375
20.5469 | -35.2139 | -31.6443
20.553 | -35.2144 | -31.6509
20.5592 | -35.2149 | -31.6577
20.5653 | -35.2154 | -31.6647
20.5714 | -35.216 | -31.6708
20.5775 | -35.2165 | -31.6777
20.5836 | -35.217 | -31.685
20.5897 | -35.2177 | -31.6923
20.5958 | -35.2183 | -31.6988
20.6019 | -35.2187 | -31.7055
20.608 | -35.2193 | -31.7124
20.6141 | -35.2199 | -31.7186
20.6202 | -35.2204 | -31.7256
20.6263 | -35.221 | -31.7328
20.6324 | -35.2216 | -31.7393
20.6385 | -35.222 | -31.746
20.6446 | -35.2226 | -31.7529
20.6507 | -35.2231 | -31.7594
20.6568 | -35.2237 | -31.7662
20.6629 | -35.2243 | -31.7725
20.669 | -35.2248 | -31.7789
20.6751 | -35.2254 | -31.7854
20.6812 | -35.2259 | -31.7923
20.6873 | -35.2265 | -31.7989
20.6934 | -35.2271 | -31.8056
20.6995 | -35.2276 | -31.8123
20.7057 | -35.2282 | -31.8192
20.7118 | -35.2287 | -31.8258
20.7179 | -35.2292 | -31.832
20.724 | -35.2297 | -31.8389
20.7301 | -35.2302 | -31.846
20.7362 | -35.2307 | -31.8528
20.7423 | -35.2312 | -31.8593
20.7484 | -35.2318 | -31.8657
20.7545 | -35.2324 | -31.8722
20.7606 | -35.2329 | -31.8789
20.7667 | -35.2334 | -31.8858
20.7728 | -35.234 | -31.8918
20.7789 | -35.2345 | -31.8987
20.785 | -35.2351 | -31.9061
20.7911 | -35.2356 | -31.9133
20.7972 | -35.2361 | -31.9202
20.8033 | -35.2366 | -31.9266
20.8094 | -35.2373 | -31.9328
20.8155 | -35.2377 | -31.9396
20.8216 | -35.2383 | -31.9468
20.8277 | -35.2387 | -31.9531
20.8338 | -35.2392 | -31.9602
20.8399 | -35.2397 | -31.9669
20.8461 | -35.2403 | -31.974
20.8522 | -35.2409 | -31.9809
20.8583 | -35.2413 | -31.9872
20.8644 | -35.2418 | -31.994
20.8705 | -35.2423 | -32.001
20.8766 | -35.2429 | -32.0081
20.8827 | -35.2434 | -32.0144
20.8888 | -35.2439 | -32.0207
20.8949 | -35.2446 | -32.0278
20.901 | -35.2451 | -32.0339
20.9071 | -35.2456 | -32.0406
20.9132 | -35.2461 | -32.0472
20.9193 | -35.2468 | -32.0538
20.9254 | -35.2472 | -32.0613
20.9315 | -35.2478 | -32.0683
20.9376 | -35.2482 | -32.0751
20.9437 | -35.2487 | -32.0829
20.9498 | -35.2493 | -32.0892
20.9559 | -35.2498 | -32.0953
20.962 | -35.2504 | -32.102
20.9681 | -35.251 | -32.1084
20.9742 | -35.2515 | -32.1155
20.9803 | -35.252 | -32.1227
20.9864 | -35.2526 | -32.1294
20.9926 | -35.2532 | -32.1364
20.9987 | -35.2536 | -32.143
21.0048 | -35.2541 | -32.1505
21.0109 | -35.2547 | -32.1577
21.017 | -35.2552 | -32.1644
21.0231 | -35.2557 | -32.1711
21.0292 | -35.2562 | -32.1775
21.0353 | -35.2569 | -32.1843
21.0414 | -35.2574 | -32.1911
21.0475 | -35.258 | -32.1976
21.0536 | -35.2586 | -32.2044
21.0597 | -35.2591 | -32.2111
21.0658 | -35.2596 | -32.2177
21.0719 | -35.2602 | -32.2241
21.078 | -35.2606 | -32.2306
21.0841 | -35.2612 | -32.2372
21.0902 | -35.2617 | -32.2442
21.0963 | -35.2622 | -32.2507
21.1024 | -35.2627 | -32.2579
21.1085 | -35.2633 | -32.2638
21.1146 | -35.2638 | -32.2709
21.1207 | -35.2644 | -32.277
21.1268 | -35.2649 | -32.2837
21.133 | -35.2654 | -32.2903
21.1391 | -35.266 | -32.2972
21.1452 | -35.2666 | -32.3035
21.1513 | -35.2671 | -32.3104
21.1574 | -35.2676 | -32.3176
21.1635 | -35.2682 | -32.3235
21.1696 | -35.2687 | -32.3299
21.1757 | -35.269 | -32.3376
21.1818 | -35.2695 | -32.3448
21.1879 | -35.2698 | -32.3511
21.194 | -35.2704 | -32.3576
21.2001 | -35.2709 | -32.364
21.2062 | -35.2714 | -32.3704
21.2123 | -35.272 | -32.3761
21.2184 | -35.2724 | -32.3826
21.2245 | -35.2729 | -32.3902
21.2306 | -35.2735 | -32.3963
21.2367 | -35.274 | -32.4028
21.2428 | -35.2746 | -32.409
21.2489 | -35.2753 | -32.4154
21.255 | -35.2758 | -32.4225
21.2611 | -35.2764 | -32.4293
21.2672 | -35.277 | -32.4363
21.2733 | -35.2775 | -32.4426
21.2795 | -35.2782 | -32.4479
21.2856 | -35.2787 | -32.4547
21.2917 | -35.2792 | -32.4618
21.2978 | -35.2797 | -32.4686
21.3039 | -35.2801 | -32.4744
21.31 | -35.2806 | -32.4808
21.3161 | -35.2813 | -32.4867
21.3222 | -35.2819 | -32.493
21.3283 | -35.2822 | -32.5
21.3344 | -35.2827 | -32.5073
21.3405 | -35.2831 | -32.5144
21.3466 | -35.2837 | -32.5217
21.3527 | -35.2841 | -32.5279
21.3588 | -35.2847 | -32.5348
21.3649 | -35.2854 | -32.5416
21.371 | -35.2861 | -32.548
21.3771 | -35.2866 | -32.5541
21.3832 | -35.2871 | -32.5608
21.3893 | -35.2877 | -32.5678
21.3954 | -35.2882 | -32.5751
21.4015 | -35.2888 | -32.5819
21.4076 | -35.2893 | -32.5882
21.4137 | -35.2898 | -32.5951
21.4199 | -35.2903 | -32.6023
21.426 | -35.2909 | -32.608
21.4321 | -35.2914 | -32.6139
21.4382 | -35.292 | -32.6202
21.4443 | -35.2925 | -32.6274
21.4504 | -35.293 | -32.6342
21.4565 | -35.2936 | -32.6402
21.4626 | -35.2941 | -32.6465
21.4687 | -35.2947 | -32.6533
21.4748 | -35.2953 | -32.6595
21.4809 | -35.2958 | -32.667
21.487 | -35.2964 | -32.6741
21.4931 | -35.297 | -32.6805
21.4992 | -35.2974 | -32.6868
21.5053 | -35.298 | -32.6938
21.5114 | -35.2987 | -32.7007
21.5175 | -35.2991 | -32.7074
21.5236 | -35.2995 | -32.7145
21.5297 | -35.3001 | -32.7216
21.5358 | -35.3007 | -32.7279
21.5419 | -35.3013 | -32.7347
21.548 | -35.3017 | -32.7409
21.5541 | -35.3023 | -32.7471
21.5602 | -35.3028 | -32.7539
21.5664 | -35.3034 | -32.7592
21.5725 | -35.3039 | -32.7657
21.5786 | -35.3044 | -32.7717
21.5847 | -35.3048 | -32.778
21.5908 | -35.3053 | -32.7847
21.5969 | -35.3058 | -32.7906
21.603 | -35.3063 | -32.7971
21.6091 | -35.3068 | -32.8036
21.6152 | -35.3075 | -32.811
21.6213 | -35.3079 | -32.8176
21.6274 | -35.3085 | -32.8239
21.6335 | -35.3091 | -32.8301
21.6396 | -35.3096 | -32.8364
21.6457 | -35.3101 | -32.8435
21.6518 | -35.3106 | -32.8499
21.6579 | -35.3111 | -32.8561
21.664 | -35.3115 | -32.8629
21.6701 | -35.3122 | -32.8697
21.6762 | -35.3127 | -32.8761
21.6823 | -35.3132 | -32.8823
21.6884 | -35.3136 | -32.8886
21.6945 | -35.3141 | -32.8947
21.7006 | -35.3148 | -32.9018
21.7068 | -35.3153 | -32.9087
21.7129 | -35.3158 | -32.9156
21.719 | -35.3164 | -32.9222
21.7251 | -35.3169 | -32.9286
21.7312 | -35.3175 | -32.934
21.7373 | -35.3183 | -32.9403
21.7434 | -35.3189 | -32.947
21.7495 | -35.3196 | -32.9534
21.7556 | -35.3201 | -32.96
21.7617 | -35.3207 | -32.9663
21.7678 | -35.3213 | -32.9735
21.7739 | -35.3217 | -32.9799
21.78 | -35.3222 | -32.9868
21.7861 | -35.3228 | -32.9941
21.7922 | -35.3234 | -33.0006
21.7983 | -35.3239 | -33.0072
21.8044 | -35.3246 | -33.0141
21.8105 | -35.3253 | -33.0213
21.8166 | -35.3258 | -33.0275
21.8227 | -35.3263 | -33.0337
21.8288 | -35.3268 | -33.04
21.8349 | -35.3275 | -33.0464
21.841 | -35.3281 | -33.0532
21.8471 | -35.3286 | -33.0598
21.8533 | -35.3291 | -33.0665
21.8594 | -35.3296 | -33.0735
21.8655 | -35.3301 | -33.0809
21.8716 | -35.3306 | -33.0867
21.8777 | -35.3311 | -33.0936
21.8838 | -35.3317 | -33.1002
21.8899 | -35.3321 | -33.1067
21.896 | -35.3326 | -33.1137
21.9021 | -35.3332 | -33.1202
21.9082 | -35.3337 | -33.1267
21.9143 | -35.3342 | -33.1325
21.9204 | -35.3348 | -33.1398
21.9265 | -35.3353 | -33.1463
21.9326 | -35.3358 | -33.1528
21.9387 | -35.3364 | -33.159
21.9448 | -35.337 | -33.1658
21.9509 | -35.3373 | -33.1722
21.957 | -35.3379 | -33.1787
21.9631 | -35.3384 | -33.1859
21.9692 | -35.3389 | -33.1929
21.9753 | -35.3395 | -33.1995
21.9814 | -35.3401 | -33.2061
21.9875 | -35.3406 | -33.2121
21.9937 | -35.341 | -33.2183
21.9998 | -35.3416 | -33.2252
22.0059 | -35.3421 | -33.2316
22.012 | -35.3427 | -33.2379
22.0181 | -35.3433 | -33.2443
22.0242 | -35.3439 | -33.2507
22.0303 | -35.3444 | -33.2571
22.0364 | -35.345 | -33.2633
22.0425 | -35.3456 | -33.2708
22.0486 | -35.3461 | -33.2765
22.0547 | -35.3466 | -33.2843
22.0608 | -35.347 | -33.2904
22.0669 | -35.3474 | -33.297
22.073 | -35.3479 | -33.3031
22.0791 | -35.3484 | -33.3093
22.0852 | -35.3489 | -33.3153
22.0913 | -35.3494 | -33.321
22.0974 | -35.3499 | -33.3267
22.1035 | -35.3503 | -33.333
22.1096 | -35.3507 | -33.3403
22.1157 | -35.3513 | -33.347
22.1218 | -35.3519 | -33.3525
22.1279 | -35.3524 | -33.3595
22.134 | -35.3529 | -33.3661
22.1402 | -35.3535 | -33.3729
22.1463 | -35.354 | -33.3791
22.1524 | -35.3545 | -33.386
22.1585 | -35.3551 | -33.3932
22.1646 | -35.3557 | -33.3994
22.1707 | -35.3561 | -33.4068
22.1768 | -35.3567 | -33.413
22.1829 | -35.3573 | -33.42
22.189 | -35.3578 | -33.4259
22.1951 | -35.3582 | -33.4321
22.2012 | -35.3589 | -33.4398
22.2073 | -35.3594 | -33.4466
22.2134 | -35.3599 | -33.4534
22.2195 | -35.3605 | -33.4597
22.2256 | -35.361 | -33.4659
22.2317 | -35.3615 | -33.4726
22.2378 | -35.362 | -33.4795
22.2439 | -35.3625 | -33.4863
22.25 | -35.3631 | -33.4934
22.2561 | -35.3636 | -33.4996
22.2622 | -35.3642 | -33.5058
22.2683 | -35.3646 | -33.512
22.2744 | -35.3651 | -33.518
22.2806 | -35.3656 | -33.5247
22.2867 | -35.3661 | -33.5316
22.2928 | -35.3667 | -33.5382
22.2989 | -35.3672 | -33.5444
22.305 | -35.3676 | -33.5513
22.3111 | -35.3683 | -33.5582
22.3172 | -35.3688 | -33.5645
22.3233 | -35.3695 | -33.5706
22.3294 | -35.3702 | -33.5771
22.3355 | -35.3708 | -33.584
22.3416 | -35.3713 | -33.5903
22.3477 | -35.3718 | -33.5969
22.3538 | -35.3724 | -33.6029
22.3599 | -35.3729 | -33.6092
22.366 | -35.3736 | -33.6155
22.3721 | -35.374 | -33.6217
22.3782 | -35.3745 | -33.6279
22.3843 | -35.3751 | -33.6347
22.3904 | -35.3757 | -33.6412
22.3965 | -35.3763 | -33.6478
22.4026 | -35.3768 | -33.6543
22.4087 | -35.3773 | -33.6603
22.4148 | -35.3779 | -33.666
22.4209 | -35.3785 | -33.6732
22.4271 | -35.379 | -33.6804
22.4332 | -35.3796 | -33.6862
22.4393 | -35.3802 | -33.6919
22.4454 | -35.3807 | -33.6982
22.4515 | -35.3811 | -33.7039
22.4576 | -35.3818 | -33.7105
22.4637 | -35.3823 | -33.7174
22.4698 | -35.3829 | -33.7246
22.4759 | -35.3836 | -33.7308
22.482 | -35.3842 | -33.7371
22.4881 | -35.3848 | -33.7441
22.4942 | -35.3853 | -33.7503
22.5003 | -35.3858 | -33.7572
22.5064 | -35.3862 | -33.763
22.5125 | -35.3869 | -33.7701
22.5186 | -35.3874 | -33.7767
22.5247 | -35.388 | -33.7834
22.5308 | -35.3885 | -33.7898
22.5369 | -35.3891 | -33.7964
22.543 | -35.3896 | -33.8022
22.5491 | -35.3902 | -33.8081
22.5552 | -35.3907 | -33.8144
22.5613 | -35.3913 | -33.8211
22.5675 | -35.3918 | -33.8278
22.5736 | -35.3924 | -33.8342
22.5797 | -35.3929 | -33.8402
22.5858 | -35.3935 | -33.8466
22.5919 | -35.3941 | -33.8526
22.598 | -35.3947 | -33.8591
22.6041 | -35.3952 | -33.8656
22.6102 | -35.3957 | -33.8725
22.6163 | -35.3963 | -33.8783
22.6224 | -35.3969 | -33.8847
22.6285 | -35.3975 | -33.8912
22.6346 | -35.398 | -33.8981
22.6407 | -35.3986 | -33.9043
22.6468 | -35.3991 | -33.9113
22.6529 | -35.3996 | -33.9177
22.659 | -35.4001 | -33.9246
22.6651 | -35.4005 | -33.9309
22.6712 | -35.4011 | -33.9371
22.6773 | -35.4015 | -33.9436
22.6834 | -35.4021 | -33.9496
22.6895 | -35.4027 | -33.9565
22.6956 | -35.4032 | -33.963
22.7017 | -35.4037 | -33.9693
22.7079 | -35.4042 | -33.9761
22.714 | -35.4047 | -33.9821
22.7201 | -35.4052 | -33.9889
22.7262 | -35.4057 | -33.9954
22.7323 | -35.4063 | -34.0018
22.7384 | -35.4068 | -34.0084
22.7445 | -35.4075 | -34.0151
22.7506 | -35.408 | -34.0219
22.7567 | -35.4085 | -34.0286
22.7628 | -35.4091 | -34.0354
22.7689 | -35.4096 | -34.0416
22.775 | -35.4101 | -34.0485
22.7811 | -35.4108 | -34.0551
22.7872 | -35.4113 | -34.0615
22.7933 | -35.4117 | -34.0686
22.7994 | -35.4122 | -34.0751
22.8055 | -35.4128 | -34.0817
22.8116 | -35.4133 | -34.0882
22.8177 | -35.4139 | -34.0945
22.8238 | -35.4143 | -34.1007
22.8299 | -35.415 | -34.107
22.836 | -35.4156 | -34.1139
22.8421 | -35.4161 | -34.1204
22.8482 | -35.4166 | -34.1269
22.8544 | -35.4173 | -34.1335
22.8605 | -35.4178 | -34.1394
22.8666 | -35.4183 | -34.1455
22.8727 | -35.4188 | -34.1523
22.8788 | -35.4194 | -34.1589
22.8849 | -35.42 | -34.1659
22.891 | -35.4205 | -34.1727
22.8971 | -35.4211 | -34.1798
22.9032 | -35.4215 | -34.1855
22.9093 | -35.422 | -34.1926
22.9154 | -35.4225 | -34.1984
22.9215 | -35.423 | -34.205
22.9276 | -35.4236 | -34.2114
22.9337 | -35.4242 | -34.2177
22.9398 | -35.4248 | -34.2244
22.9459 | -35.4255 | -34.2311
22.952 | -35.4261 | -34.2376
22.9581 | -35.4266 | -34.2437
22.9642 | -35.4272 | -34.2509
22.9703 | -35.4277 | -34.2569
22.9764 | -35.4282 | -34.2639
22.9825 | -35.4287 | -34.2707
22.9886 | -35.4293 | -34.277
22.9948 | -35.4299 | -34.2839
23.0009 | -35.4304 | -34.2903
23.007 | -35.4311 | -34.2965
23.0131 | -35.4316 | -34.3033
23.0192 | -35.4322 | -34.3097
23.0253 | -35.4327 | -34.3164
23.0314 | -35.4332 | -34.3222
23.0375 | -35.4339 | -34.3291
23.0436 | -35.4344 | -34.335
23.0497 | -35.435 | -34.3418
23.0558 | -35.4356 | -34.3485
23.0619 | -35.4361 | -34.3548
23.068 | -35.4366 | -34.3617
23.0741 | -35.4372 | -34.3677
23.0802 | -35.4377 | -34.3744
23.0863 | -35.4382 | -34.3808
23.0924 | -35.4387 | -34.3876
23.0985 | -35.4392 | -34.3941
23.1046 | -35.4397 | -34.401
23.1107 | -35.4403 | -34.407
23.1168 | -35.4409 | -34.4139
23.1229 | -35.4414 | -34.4203
23.129 | -35.442 | -34.4263
23.1351 | -35.4425 | -34.4329
23.1413 | -35.4431 | -34.439
23.1474 | -35.4436 | -34.446
23.1535 | -35.4441 | -34.4522
23.1596 | -35.4446 | -34.4581
23.1657 | -35.4452 | -34.4645
23.1718 | -35.4457 | -34.4708
23.1779 | -35.4463 | -34.4781
23.184 | -35.4469 | -34.4848
23.1901 | -35.4474 | -34.4918
23.1962 | -35.448 | -34.4981
23.2023 | -35.4484 | -34.5042
23.2084 | -35.449 | -34.5105
23.2145 | -35.4495 | -34.5167
23.2206 | -35.45 | -34.5233
23.2267 | -35.4506 | -34.5296
23.2328 | -35.4511 | -34.5361
23.2389 | -35.4516 | -34.5421
23.245 | -35.4521 | -34.5484
23.2511 | -35.4526 | -34.5552
23.2572 | -35.4532 | -34.5618
23.2633 | -35.4538 | -34.5683
23.2694 | -35.4544 | -34.5751
23.2755 | -35.4549 | -34.5819
23.2817 | -35.4554 | -34.5884
23.2878 | -35.4561 | -34.5946
23.2939 | -35.4567 | -34.6015
23.3 | -35.4573 | -34.6079
23.3061 | -35.4579 | -34.6146
23.3122 | -35.4584 | -34.6211
23.3183 | -35.4589 | -34.6271
23.3244 | -35.4595 | -34.6332
23.3305 | -35.46 | -34.6398
23.3366 | -35.4606 | -34.6466
23.3427 | -35.4611 | -34.6528
23.3488 | -35.4617 | -34.6595
23.3549 | -35.4622 | -34.6659
23.361 | -35.4628 | -34.6718
23.3671 | -35.4633 | -34.678
23.3732 | -35.4639 | -34.6844
23.3793 | -35.4644 | -34.6906
23.3854 | -35.4649 | -34.6969
23.3915 | -35.4654 | -34.7031
23.3976 | -35.4659 | -34.7102
23.4037 | -35.4664 | -34.7164
23.4098 | -35.4669 | -34.7234
23.4159 | -35.4675 | -34.7306
23.422 | -35.4681 | -34.7363
23.4282 | -35.4687 | -34.7428
23.4343 | -35.4692 | -34.7489
23.4404 | -35.4698 | -34.7558
23.4465 | -35.4702 | -34.7623
23.4526 | -35.4707 | -34.7683
23.4587 | -35.4713 | -34.7739
23.4648 | -35.4719 | -34.7808
23.4709 | -35.4723 | -34.787
23.477 | -35.4729 | -34.7935
23.4831 | -35.4735 | -34.7992
23.4892 | -35.474 | -34.805
23.4953 | -35.4746 | -34.8115
23.5014 | -35.4751 | -34.818
23.5075 | -35.4758 | -34.8243
23.5136 | -35.4762 | -34.8306
23.5197 | -35.4767 | -34.8371
23.5258 | -35.4772 | -34.8438
23.5319 | -35.4778 | -34.8509
23.538 | -35.4782 | -34.8578
23.5441 | -35.4787 | -34.8646
23.5502 | -35.4792 | -34.871
23.5563 | -35.4798 | -34.8771
23.5624 | -35.4803 | -34.8833
23.5686 | -35.4809 | -34.8888
23.5747 | -35.4814 | -34.8948
23.5808 | -35.4819 | -34.9014
23.5869 | -35.4824 | -34.9071
23.593 | -35.4829 | -34.9134
23.5991 | -35.4835 | -34.9201
23.6052 | -35.4841 | -34.9252
23.6113 | -35.4846 | -34.9315
23.6174 | -35.4852 | -34.9375
23.6235 | -35.4857 | -34.9437
23.6296 | -35.4862 | -34.9504
23.6357 | -35.4867 | -34.9565
23.6418 | -35.4874 | -34.9629
23.6479 | -35.488 | -34.9687
23.654 | -35.4885 | -34.9753
23.6601 | -35.489 | -34.9814
23.6662 | -35.4895 | -34.9881
23.6723 | -35.49 | -34.994
23.6784 | -35.4906 | -35.0003
23.6845 | -35.4911 | -35.0072
23.6906 | -35.4916 | -35.014
23.6967 | -35.4921 | -35.0208
23.7028 | -35.4926 | -35.027
23.7089 | -35.4933 | -35.0338
23.7151 | -35.4937 | -35.04
23.7212 | -35.4943 | -35.0466
23.7273 | -35.4949 | -35.0527
23.7334 | -35.4955 | -35.0589
23.7395 | -35.496 | -35.0644
23.7456 | -35.4962 | -35.0697
23.7517 | -35.4968 | -35.076
23.7578 | -35.4974 | -35.0828
23.7639 | -35.498 | -35.089
23.77 | -35.4985 | -35.096
23.7761 | -35.4991 | -35.1026
23.7822 | -35.4997 | -35.1084
23.7883 | -35.5002 | -35.1141
23.7944 | -35.5008 | -35.1199
23.8005 | -35.5013 | -35.1269
23.8066 | -35.5018 | -35.1332
23.8127 | -35.5023 | -35.1397
23.8188 | -35.5029 | -35.1463
23.8249 | -35.5034 | -35.1532
23.831 | -35.5039 | -35.159
23.8371 | -35.5044 | -35.1652
23.8432 | -35.5049 | -35.1717
23.8493 | -35.5055 | -35.1785
23.8555 | -35.506 | -35.1849
23.8616 | -35.5065 | -35.1915
23.8677 | -35.507 | -35.1978
23.8738 | -35.5074 | -35.2044
23.8799 | -35.5079 | -35.2104
23.886 | -35.5085 | -35.2172
23.8921 | -35.5091 | -35.2235
23.8982 | -35.5096 | -35.2298
23.9043 | -35.5101 | -35.236
23.9104 | -35.5106 | -35.2423
23.9165 | -35.5111 | -35.2485
23.9226 | -35.5116 | -35.254
23.9287 | -35.5122 | -35.2604
23.9348 | -35.5127 | -35.2662
23.9409 | -35.5132 | -35.2728
23.947 | -35.5137 | -35.2785
23.9531 | -35.5142 | -35.2842
23.9592 | -35.5147 | -35.2904
23.9653 | -35.5152 | -35.2962
23.9714 | -35.5159 | -35.3026
23.9775 | -35.5164 | -35.309
23.9836 | -35.517 | -35.3153
23.9897 | -35.5176 | -35.3214
23.9958 | -35.518 | -35.3274
24.002 | -35.5185 | -35.3332
24.0081 | -35.519 | -35.3387
24.0142 | -35.5194 | -35.3447
24.0203 | -35.5199 | -35.3507
24.0264 | -35.5204 | -35.3567
24.0325 | -35.521 | -35.3629
24.0386 | -35.5214 | -35.3684
24.0447 | -35.5219 | -35.3741
24.0508 | -35.5224 | -35.3801
24.0569 | -35.5229 | -35.3867
24.063 | -35.5234 | -35.3927
24.0691 | -35.5238 | -35.3984
24.0752 | -35.5244 | -35.4052
24.0813 | -35.525 | -35.4117
24.0874 | -35.5254 | -35.4183
24.0935 | -35.5259 | -35.4246
24.0996 | -35.5263 | -35.4315
24.1057 | -35.5268 | -35.4378
24.1118 | -35.5273 | -35.4442
24.1179 | -35.5279 | -35.4503
24.124 | -35.5285 | -35.4566
24.1301 | -35.5291 | -35.463
24.1362 | -35.5296 | -35.4697
24.1424 | -35.5301 | -35.4751
24.1485 | -35.5307 | -35.4806
24.1546 | -35.5314 | -35.4871
24.1607 | -35.5319 | -35.4934
24.1668 | -35.5323 | -35.4994
24.1729 | -35.5329 | -35.5058
24.179 | -35.5334 | -35.5128
24.1851 | -35.5338 | -35.5192
24.1912 | -35.5344 | -35.5264
24.1973 | -35.535 | -35.5325
24.2034 | -35.5355 | -35.5391
24.2095 | -35.536 | -35.5461
24.2156 | -35.5365 | -35.5524
24.2217 | -35.537 | -35.5584
24.2278 | -35.5376 | -35.5634
24.2339 | -35.5381 | -35.5694
24.24 | -35.5388 | -35.5754
24.2461 | -35.5393 | -35.5815
24.2522 | -35.5399 | -35.5878
24.2583 | -35.5404 | -35.5945
24.2644 | -35.541 | -35.6004
24.2705 | -35.5415 | -35.6072
24.2766 | -35.542 | -35.6133
24.2827 | -35.5426 | -35.6195
24.2889 | -35.5432 | -35.625
24.295 | -35.5436 | -35.6317
24.3011 | -35.5442 | -35.6381
24.3072 | -35.5447 | -35.6446
24.3133 | -35.5451 | -35.6508
24.3194 | -35.5456 | -35.6577
24.3255 | -35.5463 | -35.6639
24.3316 | -35.5468 | -35.6695
24.3377 | -35.5473 | -35.6761
24.3438 | -35.5478 | -35.6818
24.3499 | -35.5483 | -35.6882
24.356 | -35.5487 | -35.6943
24.3621 | -35.5492 | -35.7007
24.3682 | -35.5497 | -35.7065
24.3743 | -35.5503 | -35.7132
24.3804 | -35.5507 | -35.7179
24.3865 | -35.5511 | -35.7244
24.3926 | -35.5517 | -35.7303
24.3987 | -35.5523 | -35.7367
24.4048 | -35.5528 | -35.7425
24.4109 | -35.5533 | -35.7477
24.417 | -35.5538 | -35.7536
24.4231 | -35.5544 | -35.76
24.4293 | -35.5549 | -35.7664
24.4354 | -35.5554 | -35.7727
24.4415 | -35.556 | -35.7787
24.4476 | -35.5566 | -35.7843
24.4537 | -35.5572 | -35.7906
24.4598 | -35.5578 | -35.7966
24.4659 | -35.5582 | -35.8028
24.472 | -35.5588 | -35.8091
24.4781 | -35.5594 | -35.8152
24.4842 | -35.5598 | -35.8209
24.4903 | -35.5604 | -35.8263
24.4964 | -35.5609 | -35.8326
24.5025 | -35.5614 | -35.8391
24.5086 | -35.562 | -35.8454
24.5147 | -35.5625 | -35.8515
24.5208 | -35.563 | -35.8572
24.5269 | -35.5635 | -35.8632
24.533 | -35.5641 | -35.8688
24.5391 | -35.5646 | -35.8748
24.5452 | -35.5653 | -35.8801
24.5513 | -35.566 | -35.8861
24.5574 | -35.5664 | -35.8923
24.5635 | -35.567 | -35.8986
24.5696 | -35.5675 | -35.9043
24.5758 | -35.5681 | -35.9105
24.5819 | -35.5685 | -35.9165
24.588 | -35.5691 | -35.9217
24.5941 | -35.5697 | -35.9271
24.6002 | -35.5702 | -35.9331
24.6063 | -35.5707 | -35.9394
24.6124 | -35.5713 | -35.9457
24.6185 | -35.5719 | -35.9516
24.6246 | -35.5725 | -35.9577
24.6307 | -35.5728 | -35.964
24.6368 | -35.5734 | -35.97
24.6429 | -35.574 | -35.9766
24.649 | -35.5745 | -35.9826
24.6551 | -35.575 | -35.9888
24.6612 | -35.5754 | -35.9949
24.6673 | -35.5758 | -36.0009
24.6734 | -35.5764 | -36.0071
24.6795 | -35.5769 | -36.0135
24.6856 | -35.5773 | -36.0193
24.6917 | -35.5778 | -36.0253
24.6978 | -35.5784 | -36.0302
24.7039 | -35.5789 | -36.0358
24.71 | -35.5795 | -36.0412
24.7162 | -35.5799 | -36.0479
24.7223 | -35.5803 | -36.0537
24.7284 | -35.5808 | -36.0586
24.7345 | -35.5814 | -36.0641
24.7406 | -35.582 | -36.0697
24.7467 | -35.5826 | -36.0757
24.7528 | -35.583 | -36.0814
24.7589 | -35.5835 | -36.0868
24.765 | -35.5841 | -36.0928
24.7711 | -35.5845 | -36.0986
24.7772 | -35.585 | -36.1046
24.7833 | -35.5855 | -36.1107
24.7894 | -35.586 | -36.1164
24.7955 | -35.5865 | -36.1225
24.8016 | -35.587 | -36.1279
24.8077 | -35.5876 | -36.1345
24.8138 | -35.5882 | -36.1406
24.8199 | -35.5887 | -36.1471
24.826 | -35.5892 | -36.1528
24.8321 | -35.5897 | -36.1592
24.8382 | -35.5901 | -36.1648
24.8443 | -35.5906 | -36.1708
24.8504 | -35.591 | -36.1767
24.8565 | -35.5916 | -36.1828
24.8627 | -35.5921 | -36.1885
24.8688 | -35.5927 | -36.1922
24.8749 | -35.5933 | -36.1983
24.881 | -35.5937 | -36.2042
24.8871 | -35.5942 | -36.2098
24.8932 | -35.5947 | -36.2157
24.8993 | -35.5952 | -36.2212
24.9054 | -35.5957 | -36.2275
24.9115 | -35.5962 | -36.2343
24.9176 | -35.5966 | -36.2402
24.9237 | -35.5971 | -36.2459
24.9298 | -35.5977 | -36.2521
24.9359 | -35.5982 | -36.2578
24.942 | -35.5987 | -36.2636
24.9481 | -35.5992 | -36.27
24.9542 | -35.5996 | -36.276
24.9603 | -35.6001 | -36.2824
24.9664 | -35.6004 | -36.2879
24.9725 | -35.6009 | -36.2938
24.9786 | -35.6013 | -36.2997
24.9847 | -35.6017 | -36.3059
24.9908 | -35.6023 | -36.3117
24.9969 | -35.6028 | -36.317
25.0031 | -35.6034 | -36.3229
25.0092 | -35.6038 | -36.3273
25.0153 | -35.6044 | -36.3329
25.0214 | -35.6049 | -36.3384
25.0275 | -35.6056 | -36.3447
25.0336 | -35.6061 | -36.3507
25.0397 | -35.6066 | -36.3571
25.0458 | -35.607 | -36.362
25.0519 | -35.6075 | -36.3684
25.058 | -35.6079 | -36.3743
25.0641 | -35.6085 | -36.3804
25.0702 | -35.6091 | -36.3864
25.0763 | -35.6095 | -36.3918
25.0824 | -35.61 | -36.3963
25.0885 | -35.6104 | -36.4024
25.0946 | -35.6109 | -36.4083
25.1007 | -35.6115 | -36.415
25.1068 | -35.6119 | -36.4203
25.1129 | -35.6124 | -36.4258
25.119 | -35.6129 | -36.4306
25.1251 | -35.6133 | -36.4364
25.1312 | -35.6139 | -36.4419
25.1373 | -35.6143 | -36.4473
25.1435 | -35.6149 | -36.4531
25.1496 | -35.6154 | -36.4583
25.1557 | -35.616 | -36.4634
25.1618 | -35.6166 | -36.4693
25.1679 | -35.6172 | -36.4752
25.174 | -35.6176 | -36.4806
25.1801 | -35.6182 | -36.4867
25.1862 | -35.6185 | -36.4919
25.1923 | -35.619 | -36.4977
25.1984 | -35.6195 | -36.5032
25.2045 | -35.6199 | -36.5089
25.2106 | -35.6203 | -36.5145
25.2167 | -35.6209 | -36.5198
25.2228 | -35.6213 | -36.5244
25.2289 | -35.6218 | -36.5299
25.235 | -35.6222 | -36.5358
25.2411 | -35.6227 | -36.5419
25.2472 | -35.6232 | -36.5471
25.2533 | -35.6237 | -36.5533
25.2594 | -35.6241 | -36.5581
25.2655 | -35.6245 | -36.5637
25.2716 | -35.6251 | -36.5692
25.2777 | -35.6257 | -36.5744
25.2838 | -35.6261 | -36.5796
25.29 | -35.6267 | -36.5854
25.2961 | -35.6271 | -36.5919
25.3022 | -35.6276 | -36.5973
25.3083 | -35.6281 | -36.6031
25.3144 | -35.6287 | -36.6084
25.3205 | -35.6291 | -36.6138
25.3266 | -35.6295 | -36.618
25.3327 | -35.63 | -36.6239
25.3388 | -35.6304 | -36.6296
25.3449 | -35.6308 | -36.6359
25.351 | -35.6313 | -36.6407
25.3571 | -35.6317 | -36.6467
25.3632 | -35.6322 | -36.6511
25.3693 | -35.6326 | -36.6573
25.3754 | -35.6332 | -36.6631
25.3815 | -35.6335 | -36.6693
25.3876 | -35.6341 | -36.6741
25.3937 | -35.6348 | -36.6794
25.3998 | -35.6352 | -36.6839
25.4059 | -35.6356 | -36.6895
25.412 | -35.6363 | -36.695
25.4181 | -35.6367 | -36.7011
25.4242 | -35.6372 | -36.7065
25.4304 | -35.6377 | -36.7116
25.4365 | -35.6383 | -36.7167
25.4426 | -35.6388 | -36.7227
25.4487 | -35.6393 | -36.728
25.4548 | -35.6397 | -36.7342
25.4609 | -35.6403 | -36.7399
25.467 | -35.641 | -36.7465
25.4731 | -35.6415 | -36.7523
25.4792 | -35.6419 | -36.7578
25.4853 | -35.6426 | -36.7633
25.4914 | -35.6431 | -36.7697
25.4975 | -35.6436 | -36.775
25.5036 | -35.6442 | -36.7798
25.5097 | -35.6446 | -36.786
25.5158 | -35.6453 | -36.7916
25.5219 | -35.6458 | -36.7973
25.528 | -35.6462 | -36.802
25.5341 | -35.6467 | -36.809
25.5402 | -35.6471 | -36.8146
25.5463 | -35.6476 | -36.8189
25.5524 | -35.6482 | -36.8252
25.5585 | -35.6486 | -36.8292
25.5646 | -35.6493 | -36.835
25.5707 | -35.6499 | -36.8402
25.5769 | -35.6506 | -36.8468
25.583 | -35.6511 | -36.8518
25.5891 | -35.6517 | -36.8565
25.5952 | -35.6522 | -36.8623
25.6013 | -35.6526 | -36.8675
25.6074 | -35.6532 | -36.8734
25.6135 | -35.6538 | -36.8782
25.6196 | -35.6542 | -36.8844
25.6257 | -35.6546 | -36.8896
25.6318 | -35.6551 | -36.8949
25.6379 | -35.6556 | -36.9007
25.644 | -35.6562 | -36.9059
25.6501 | -35.6568 | -36.9114
25.6562 | -35.6574 | -36.9173
25.6623 | -35.6578 | -36.9226
25.6684 | -35.6582 | -36.9293
25.6745 | -35.6587 | -36.9346
25.6806 | -35.6592 | -36.9398
25.6867 | -35.6597 | -36.9459
25.6928 | -35.6602 | -36.9515
25.6989 | -35.6607 | -36.9571
25.705 | -35.6611 | -36.9627
25.7111 | -35.6617 | -36.9686
25.7173 | -35.662 | -36.9723
25.7234 | -35.6624 | -36.9785
25.7295 | -35.6627 | -36.9837
25.7356 | -35.6633 | -36.9894
25.7417 | -35.6638 | -36.9948
25.7478 | -35.6643 | -36.9992
25.7539 | -35.6646 | -37.0044
25.76 | -35.6651 | -37.01
25.7661 | -35.6655 | -37.0153
25.7722 | -35.666 | -37.0212
25.7783 | -35.6665 | -37.0268
25.7844 | -35.6669 | -37.0327
25.7905 | -35.6674 | -37.0379
25.7966 | -35.6679 | -37.0428
25.8027 | -35.6685 | -37.0484
25.8088 | -35.6689 | -37.0541
25.8149 | -35.6694 | -37.0592
25.821 | -35.6698 | -37.0649
25.8271 | -35.6703 | -37.0704
25.8332 | -35.6707 | -37.0766
25.8393 | -35.6713 | -37.0824
25.8454 | -35.6716 | -37.0873
25.8515 | -35.6722 | -37.0926
25.8576 | -35.6726 | -37.0964
25.8638 | -35.6732 | -37.1025
25.8699 | -35.6737 | -37.1082
25.876 | -35.6742 | -37.1139
25.8821 | -35.6747 | -37.1194
25.8882 | -35.6752 | -37.1238
25.8943 | -35.6756 | -37.129
25.9004 | -35.676 | -37.1342
25.9065 | -35.6765 | -37.1394
25.9126 | -35.677 | -37.1454
25.9187 | -35.6774 | -37.15
25.9248 | -35.6778 | -37.1534
25.9309 | -35.6782 | -37.1585
25.937 | -35.6788 | -37.1636
25.9431 | -35.6792 | -37.1692
25.9492 | -35.6797 | -37.1741
25.9553 | -35.6801 | -37.1789
25.9614 | -35.6806 | -37.184
25.9675 | -35.681 | -37.1894
25.9736 | -35.6816 | -37.1955
25.9797 | -35.6821 | -37.2004
25.9858 | -35.6827 | -37.2052
25.9919 | -35.683 | -37.2109
25.998 | -35.6834 | -37.2161
26.0042 | -35.6839 | -37.2218
26.0103 | -35.6843 | -37.2269
26.0164 | -35.6848 | -37.2325
26.0225 | -35.6852 | -37.2384
26.0286 | -35.6858 | -37.2444
26.0347 | -35.6863 | -37.2487
26.0408 | -35.6867 | -37.2538
26.0469 | -35.6872 | -37.2589
26.053 | -35.6876 | -37.2637
26.0591 | -35.6882 | -37.2695
26.0652 | -35.6889 | -37.2746
26.0713 | -35.6893 | -37.2803
26.0774 | -35.6897 | -37.2854
26.0835 | -35.6902 | -37.2899
26.0896 | -35.6906 | -37.295
26.0957 | -35.6911 | -37.3005
26.1018 | -35.6915 | -37.3053
26.1079 | -35.6919 | -37.3107
26.114 | -35.6923 | -37.3152
26.1201 | -35.6927 | -37.32
26.1262 | -35.6931 | -37.3259
26.1323 | -35.6935 | -37.3314
26.1384 | -35.6941 | -37.3363
26.1445 | -35.6945 | -37.342
26.1507 | -35.695 | -37.3483
26.1568 | -35.6953 | -37.3528
26.1629 | -35.6957 | -37.3574
26.169 | -35.6962 | -37.362
26.1751 | -35.6967 | -37.3672
26.1812 | -35.6972 | -37.3722
26.1873 | -35.6976 | -37.3771
26.1934 | -35.6981 | -37.3818
26.1995 | -35.6985 | -37.3874
26.2056 | -35.6989 | -37.3918
26.2117 | -35.6993 | -37.3967
26.2178 | -35.6998 | -37.4017
26.2239 | -35.7003 | -37.4071
26.23 | -35.7007 | -37.4119
26.2361 | -35.7012 | -37.4175
26.2422 | -35.7014 | -37.4226
26.2483 | -35.7019 | -37.4275
26.2544 | -35.7024 | -37.4325
26.2605 | -35.7029 | -37.4378
26.2666 | -35.7033 | -37.4427
26.2727 | -35.7037 | -37.4479
26.2788 | -35.7042 | -37.4529
26.2849 | -35.7046 | -37.4573
26.2911 | -35.705 | -37.4628
26.2972 | -35.7054 | -37.4679
26.3033 | -35.7059 | -37.4736
26.3094 | -35.7063 | -37.48
26.3155 | -35.7067 | -37.485
26.3216 | -35.7071 | -37.49
26.3277 | -35.7077 | -37.4952
26.3338 | -35.7081 | -37.5007
26.3399 | -35.7086 | -37.5052
26.346 | -35.7091 | -37.5095
26.3521 | -35.7095 | -37.515
26.3582 | -35.71 | -37.5198
26.3643 | -35.7105 | -37.5255
26.3704 | -35.711 | -37.5299
26.3765 | -35.7114 | -37.5357
26.3826 | -35.7119 | -37.5405
26.3887 | -35.7124 | -37.5449
26.3948 | -35.7128 | -37.5505
26.4009 | -35.7132 | -37.5556
26.407 | -35.7136 | -37.5607
26.4131 | -35.7141 | -37.5661
26.4192 | -35.7145 | -37.5703
26.4253 | -35.7149 | -37.575
26.4314 | -35.7153 | -37.5799
26.4376 | -35.7158 | -37.5851
26.4437 | -35.7162 | -37.5907
26.4498 | -35.7167 | -37.5958
26.4559 | -35.717 | -37.6004
26.462 | -35.7176 | -37.6055
26.4681 | -35.718 | -37.6116
26.4742 | -35.7183 | -37.617
26.4803 | -35.7188 | -37.6224
26.4864 | -35.7191 | -37.627
26.4925 | -35.7197 | -37.6326
26.4986 | -35.72 | -37.6375
26.5047 | -35.7204 | -37.6427
26.5108 | -35.7209 | -37.6479
26.5169 | -35.7214 | -37.6535
26.523 | -35.7219 | -37.659
26.5291 | -35.7223 | -37.6645
26.5352 | -35.7226 | -37.6701
26.5413 | -35.7231 | -37.6755
26.5474 | -35.7234 | -37.6809
26.5535 | -35.7238 | -37.6855
26.5596 | -35.724 | -37.6902
26.5657 | -35.7245 | -37.6949
26.5718 | -35.725 | -37.6998
26.578 | -35.7255 | -37.7055
26.5841 | -35.7259 | -37.7105
26.5902 | -35.7264 | -37.716
26.5963 | -35.7267 | -37.7205
26.6024 | -35.7271 | -37.7259
26.6085 | -35.7275 | -37.73
26.6146 | -35.7278 | -37.7355
26.6207 | -35.7283 | -37.741
26.6268 | -35.7286 | -37.7463
26.6329 | -35.729 | -37.7513
26.639 | -35.7295 | -37.7568
26.6451 | -35.7299 | -37.7619
26.6512 | -35.7303 | -37.7668
26.6573 | -35.7307 | -37.7721
26.6634 | -35.7311 | -37.7768
26.6695 | -35.7315 | -37.7825
26.6756 | -35.7319 | -37.7871
26.6817 | -35.7323 | -37.7924
26.6878 | -35.7327 | -37.7972
26.6939 | -35.7333 | -37.8018
26.7 | -35.7338 | -37.8056
26.7061 | -35.7342 | -37.8108
26.7122 | -35.7346 | -37.8156
26.7183 | -35.735 | -37.8201
26.7245 | -35.7355 | -37.8254
26.7306 | -35.7358 | -37.8306
26.7367 | -35.7362 | -37.8349
26.7428 | -35.7366 | -37.84
26.7489 | -35.7371 | -37.8451
26.755 | -35.7375 | -37.85
26.7611 | -35.738 | -37.8551
26.7672 | -35.7385 | -37.86
26.7733 | -35.7388 | -37.8644
26.7794 | -35.7392 | -37.8694
26.7855 | -35.7397 | -37.8744
26.7916 | -35.7401 | -37.8787
26.7977 | -35.7406 | -37.8839
26.8038 | -35.741 | -37.8882
26.8099 | -35.7414 | -37.8935
26.816 | -35.7418 | -37.8988
26.8221 | -35.7422 | -37.904
26.8282 | -35.7427 | -37.9093
26.8343 | -35.7433 | -37.9148
26.8404 | -35.7437 | -37.9201
26.8465 | -35.7441 | -37.9245
26.8526 | -35.7446 | -37.9302
26.8587 | -35.745 | -37.9356
26.8649 | -35.7455 | -37.9403
26.871 | -35.7459 | -37.9453
26.8771 | -35.7463 | -37.9506
26.8832 | -35.7467 | -37.9554
26.8893 | -35.7471 | -37.9603
26.8954 | -35.7474 | -37.9648
26.9015 | -35.7479 | -37.9691
26.9076 | -35.7484 | -37.9743
26.9137 | -35.7487 | -37.9787
26.9198 | -35.7491 | -37.9837
26.9259 | -35.7493 | -37.9887
26.932 | -35.7497 | -37.9941
26.9381 | -35.7501 | -37.9996
26.9442 | -35.7504 | -38.0044
26.9503 | -35.7508 | -38.0099
26.9564 | -35.7511 | -38.0153
26.9625 | -35.7515 | -38.0203
26.9686 | -35.752 | -38.0251
26.9747 | -35.7523 | -38.0306
26.9808 | -35.7528 | -38.0358
26.9869 | -35.7534 | -38.04
26.993 | -35.7539 | -38.0453
26.9991 | -35.7542 | -38.0491
27.0052 | -35.7548 | -38.0539
27.0114 | -35.7552 | -38.0595
27.0175 | -35.7555 | -38.0641
27.0236 | -35.7559 | -38.0691
27.0297 | -35.7564 | -38.0746
27.0358 | -35.7567 | -38.0798
27.0419 | -35.7571 | -38.084
27.048 | -35.7575 | -38.0884
27.0541 | -35.7577 | -38.0933
27.0602 | -35.758 | -38.0991
27.0663 | -35.7585 | -38.1047
27.0724 | -35.7589 | -38.1095
27.0785 | -35.7593 | -38.1144
27.0846 | -35.7597 | -38.1193
27.0907 | -35.7602 | -38.1246
27.0968 | -35.7606 | -38.1294
27.1029 | -35.7611 | -38.1335
27.109 | -35.7614 | -38.1383
27.1151 | -35.7617 | -38.1438
27.1212 | -35.7621 | -38.1485
27.1273 | -35.7625 | -38.1533
27.1334 | -35.7628 | -38.1587
27.1395 | -35.7633 | -38.1641
27.1456 | -35.7637 | -38.1696
27.1518 | -35.7641 | -38.1744
27.1579 | -35.7644 | -38.1794
27.164 | -35.7649 | -38.1842
27.1701 | -35.7653 | -38.189
27.1762 | -35.7655 | -38.1944
27.1823 | -35.7659 | -38.1991
27.1884 | -35.7664 | -38.204
27.1945 | -35.7669 | -38.208
27.2006 | -35.7674 | -38.2124
27.2067 | -35.7677 | -38.2172
27.2128 | -35.7681 | -38.2222
27.2189 | -35.7685 | -38.2268
27.225 | -35.7689 | -38.2319
27.2311 | -35.7695 | -38.2366
27.2372 | -35.7699 | -38.2416
27.2433 | -35.7703 | -38.2466
27.2494 | -35.7707 | -38.2514
27.2555 | -35.7711 | -38.2561
27.2616 | -35.7716 | -38.2617
27.2677 | -35.7718 | -38.2664
27.2738 | -35.7723 | -38.2713
27.2799 | -35.7728 | -38.2762
27.286 | -35.7732 | -38.2816
27.2921 | -35.7737 | -38.2864
27.2983 | -35.774 | -38.2907
27.3044 | -35.7745 | -38.2952
27.3105 | -35.775 | -38.3001
27.3166 | -35.7754 | -38.3049
27.3227 | -35.7758 | -38.3091
27.3288 | -35.7762 | -38.3144
27.3349 | -35.7767 | -38.3192
27.341 | -35.777 | -38.3249
27.3471 | -35.7774 | -38.3298
27.3532 | -35.7778 | -38.3349
27.3593 | -35.7783 | -38.3406
27.3654 | -35.7788 | -38.3456
27.3715 | -35.7792 | -38.3493
27.3776 | -35.7796 | -38.3544
27.3837 | -35.7801 | -38.3587
27.3898 | -35.7805 | -38.3635
27.3959 | -35.7809 | -38.3688
27.402 | -35.7814 | -38.3732
27.4081 | -35.7818 | -38.3782
27.4142 | -35.7821 | -38.3832
27.4203 | -35.7826 | -38.3883
27.4264 | -35.7829 | -38.3933
27.4325 | -35.7834 | -38.3983
27.4387 | -35.7837 | -38.4032
27.4448 | -35.7841 | -38.4082
27.4509 | -35.7845 | -38.4133
27.457 | -35.785 | -38.4177
27.4631 | -35.7854 | -38.423
27.4692 | -35.7858 | -38.4284
27.4753 | -35.7862 | -38.4335
27.4814 | -35.7867 | -38.4386
27.4875 | -35.7871 | -38.4444
27.4936 | -35.7875 | -38.4496
27.4997 | -35.7879 | -38.4541
27.5058 | -35.7884 | -38.4589
27.5119 | -35.7887 | -38.4638
27.518 | -35.7891 | -38.469
27.5241 | -35.7895 | -38.4739
27.5302 | -35.7899 | -38.4788
27.5363 | -35.7904 | -38.4833
27.5424 | -35.7908 | -38.4878
27.5485 | -35.7913 | -38.4931
27.5546 | -35.7917 | -38.4982
27.5607 | -35.7922 | -38.5028
27.5668 | -35.7926 | -38.5074
27.5729 | -35.793 | -38.5123
27.5791 | -35.7934 | -38.517
27.5852 | -35.7938 | -38.5219
27.5913 | -35.7942 | -38.5268
27.5974 | -35.7947 | -38.5322
27.6035 | -35.795 | -38.5377
27.6096 | -35.7954 | -38.5433
27.6157 | -35.7958 | -38.548
27.6218 | -35.7962 | -38.5532
27.6279 | -35.7965 | -38.5583
27.634 | -35.7969 | -38.5639
27.6401 | -35.7973 | -38.5687
27.6462 | -35.7977 | -38.5735
27.6523 | -35.798 | -38.5781
27.6584 | -35.7984 | -38.583
27.6645 | -35.7989 | -38.5885
27.6706 | -35.7991 | -38.5933
27.6767 | -35.7996 | -38.599
27.6828 | -35.7999 | -38.6034
27.6889 | -35.8004 | -38.6087
27.695 | -35.8007 | -38.6139
27.7011 | -35.8011 | -38.6188
27.7072 | -35.8016 | -38.6237
27.7133 | -35.802 | -38.629
27.7194 | -35.8023 | -38.6336
27.7256 | -35.8026 | -38.6382
27.7317 | -35.8029 | -38.6428
27.7378 | -35.8033 | -38.6479
27.7439 | -35.8037 | -38.653
27.75 | -35.8041 | -38.6579
27.7561 | -35.8044 | -38.6637
27.7622 | -35.8047 | -38.6684
27.7683 | -35.8051 | -38.673
27.7744 | -35.8054 | -38.6778
27.7805 | -35.8058 | -38.6826
27.7866 | -35.8061 | -38.6874
27.7927 | -35.8065 | -38.6929
27.7988 | -35.807 | -38.698
27.8049 | -35.8073 | -38.7025
27.811 | -35.8077 | -38.7068
27.8171 | -35.8082 | -38.7119
27.8232 | -35.8086 | -38.7169
27.8293 | -35.809 | -38.7215
27.8354 | -35.8094 | -38.7265
27.8415 | -35.8098 | -38.7319
27.8476 | -35.8101 | -38.7365
27.8537 | -35.8105 | -38.7409
27.8598 | -35.8108 | -38.7459
27.866 | -35.8112 | -38.7508
27.8721 | -35.8116 | -38.7558
27.8782 | -35.8119 | -38.7608
27.8843 | -35.8124 | -38.7655
27.8904 | -35.8129 | -38.7704
27.8965 | -35.8133 | -38.7749
27.9026 | -35.8136 | -38.7807
27.9087 | -35.814 | -38.7861
27.9148 | -35.8143 | -38.7906
27.9209 | -35.8148 | -38.7957
27.927 | -35.8152 | -38.8009
27.9331 | -35.8157 | -38.8063
27.9392 | -35.8161 | -38.8117
27.9453 | -35.8164 | -38.8162
27.9514 | -35.8168 | -38.8213
27.9575 | -35.8171 | -38.8262
27.9636 | -35.8176 | -38.8312
27.9697 | -35.818 | -38.8363
27.9758 | -35.8185 | -38.8413
27.9819 | -35.8188 | -38.8471
27.988 | -35.8191 | -38.8519
27.9941 | -35.8195 | -38.8568
28.0002 | -35.8199 | -38.862
28.0063 | -35.8203 | -38.8672
28.0125 | -35.8206 | -38.872
28.0186 | -35.8209 | -38.8769
28.0247 | -35.8214 | -38.8819
28.0308 | -35.8219 | -38.8864
28.0369 | -35.8222 | -38.8912
28.043 | -35.8226 | -38.8965
28.0491 | -35.8229 | -38.9017
28.0552 | -35.8232 | -38.907
28.0613 | -35.8235 | -38.9118
28.0674 | -35.8238 | -38.9162
28.0735 | -35.8242 | -38.9207

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.6579 | -43.0415
21.4908 | -35.6077 | -35.6606
16.4948 | -34.9354 | -28.1534
12.6602 | -34.5464 | -21.7686
9.7171 | -34.3098 | -16.7867
7.4582 | -34.1492 | -12.7527
5.7244 | -34.0382 | -9.4857
4.3936 | -33.966 | -6.8456
3.3722 | -33.9493 | -4.7659
2.5883 | -33.9417 | -3.1895
1.9866 | -33.9361 | -1.9323
1.5248 | -33.9376 | -0.92341
1.1703 | -33.9514 | -0.26754
1 | -33.9611 | 0.05987
0.89822 | -33.954 | 0.33343
0.68941 | -33.9334 | 0.91926
0.52919 | -33.9357 | 1.3726
0.40612 | -33.939 | 1.9818
0.31171 | -33.9302 | 2.4371
0.23928 | -33.9432 | 3.1637
0.18358 | -33.9502 | 3.5801
0.14091 | -33.9483 | 4.4598
0.10819 | -33.9736 | 5.6172
0.083017 | -34.0133 | 7.0828
0.063705 | -34.0334 | 8.9561
0.048923 | -34.1038 | 11.4355
0.037479 | -34.2116 | 14.607
0.028801 | -34.3861 | 18.6008
0.022078 | -34.6692 | 23.9096
0.017023 | -35.1302 | 30.3085
0.013018 | -35.9321 | 37.9881
0.010014 | -37.012 | 45.8797

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003704 | -6.667
0.007407 | -27.1532
0.011111 | -34.8168
0.014815 | -39.2472
0.018519 | -41.1796
0.022222 | -42.7131
0.025926 | -44.9484
0.02963 | -46.2261
0.033333 | -47.1063
0.037037 | -47.9434
0.040741 | -49.4228
0.044444 | -50.013
0.048148 | -51.581
0.051852 | -52.9498
0.055556 | -53.8776
0.059259 | -54.723
0.062963 | -55.4385
0.066667 | -55.3973
0.07037 | -55.2681
0.074074 | -56.4045
0.077778 | -56.4021
0.081481 | -57.3816
0.085185 | -58.423
0.088889 | -58.9407
0.092593 | -59.4327
0.096296 | -59.1164
0.1 | -59.4549
0.1037 | -60.1394
0.10741 | -61.1146
0.11111 | -62.2536
0.11482 | -63.2088
0.11852 | -63.5043
0.12222 | -62.3508
0.12593 | -63.0664
0.12963 | -64.0462
0.13333 | -64.7325
0.13704 | -65.0261
0.14074 | -65.7854
0.14444 | -66.3103
0.14815 | -66.2984
0.15185 | -64.9076
0.15556 | -65.6441
0.15926 | -65.7638
0.16296 | -66.4798
0.16667 | -69.0046
0.17037 | -67.2495
0.17407 | -67.9463
0.17778 | -67.2964
0.18148 | -67.2999
0.18518 | -68.705
0.18889 | -69.0641
0.19259 | -68.6673
0.1963 | -68.7599
0.2 | -68.8868
0.2037 | -69.0376
0.20741 | -68.859
0.21111 | -69.21
0.21482 | -69.2856
0.21852 | -69.7932
0.22222 | -71.2063
0.22593 | -70.8389
0.22963 | -70.7144
0.23333 | -71.0002
0.23704 | -70.986
0.24074 | -71.2967
0.24444 | -70.8864
0.24815 | -71.2208
0.25185 | -72.4207
0.25556 | -71.7878
0.25926 | -71.3751
0.26296 | -71.8457
0.26667 | -72.7449
0.27037 | -73.0943
0.27407 | -72.7678
0.27778 | -72.9298
0.28148 | -72.5231
0.28519 | -73.6158
0.28889 | -74.3006
0.29259 | -73.6124
0.2963 | -74.687
0.3 | -74.8668
0.3037 | -74.4756
0.30741 | -74.6121
0.31111 | -74.1959
0.31482 | -74.1603
0.31852 | -73.7552
0.32222 | -74.1061
0.32593 | -74.9679
0.32963 | -74.9039
0.33333 | -73.8042
0.33704 | -74.4613
0.34074 | -76.8685
0.34444 | -76.6399
0.34815 | -76.1726
0.35185 | -75.5363
0.35556 | -75.8756
0.35926 | -75.711
0.36296 | -76.0315
0.36667 | -75.8153
0.37037 | -75.5006
0.37407 | -77.3894
0.37778 | -77.7243
0.38148 | -76.0823
0.38518 | -76.0201
0.38889 | -76.1294
0.39259 | -77.105
0.3963 | -76.7793
0.4 | -76.5792
0.4037 | -77.3694
0.40741 | -77.2625
0.41111 | -76.8804
0.41481 | -77.0789
0.41852 | -77.9419
0.42222 | -77.7483
0.42593 | -77.2741
0.42963 | -78.2762
0.43333 | -77.9483
0.43704 | -77.3434
0.44074 | -77.8934
0.44444 | -78.6377
0.44815 | -77.6805
0.45185 | -78.8372
0.45556 | -79.7425
0.45926 | -79.3924
0.46296 | -80.5806
0.46667 | -79.6635
0.47037 | -78.5294
0.47407 | -77.8383
0.47778 | -78.1346
0.48148 | -79.3162
0.48518 | -79.6566
0.48889 | -79.0883
0.49259 | -79.1991
0.4963 | -79.8953
0.5 | -80.207
0.5037 | -78.9452
0.50741 | -78.2829
0.51111 | -78.8731
0.51482 | -80.536
0.51852 | -80.7989
0.52222 | -79.6062
0.52593 | -79.7969
0.52963 | -80.6348
0.53333 | -80.2779
0.53704 | -80.2241
0.54074 | -80.1141
0.54444 | -79.9926
0.54815 | -80.2115
0.55185 | -80.4409
0.55556 | -80.4087
0.55926 | -80.3903
0.56296 | -80.9694
0.56667 | -81.6113
0.57037 | -81.6081
0.57407 | -81.0582
0.57778 | -80.5448
0.58148 | -80.3921
0.58518 | -80.4696
0.58889 | -81.018
0.59259 | -81.8222
0.5963 | -82.088
0.6 | -81.6537
0.6037 | -81.3078
0.60741 | -81.1566
0.61111 | -81.1685
0.61482 | -81.3259
0.61852 | -81.5669
0.62222 | -81.7404
0.62593 | -81.8328
0.62963 | -81.9286
0.63333 | -81.9326
0.63704 | -82.1254
0.64074 | -82.4242
0.64444 | -82.0835
0.64815 | -81.7365
0.65185 | -82.0459
0.65556 | -82.4494
0.65926 | -82.4081
0.66296 | -82.4608
0.66667 | -82.9134
0.67037 | -82.8349
0.67407 | -82.4115
0.67778 | -82.3456
0.68148 | -82.7606
0.68519 | -83.0441
0.68889 | -83.028
0.69259 | -83.3768
0.6963 | -83.8442
0.7 | -84.1621
0.7037 | -83.6988
0.70741 | -82.7303
0.71111 | -82.3932
0.71481 | -82.9727
0.71852 | -83.7693
0.72222 | -83.9
0.72593 | -83.5713
0.72963 | -83.282
0.73333 | -83.063
0.73704 | -83.0216
0.74074 | -83.4728
0.74444 | -83.8136
0.74815 | -83.5349
0.75185 | -83.1705
0.75556 | -83.4671
0.75926 | -83.9935
0.76296 | -84.2518
0.76667 | -84.3448
0.77037 | -83.9493
0.77407 | -83.4335
0.77778 | -83.4749
0.78148 | -83.4681
0.78519 | -83.2458
0.78889 | -83.255
0.79259 | -83.6704
0.7963 | -84.4374
0.8 | -84.9771
0.8037 | -84.7382
0.80741 | -84.2374
0.81111 | -83.9628
0.81481 | -84.1393
0.81852 | -84.5636
0.82222 | -84.4454
0.82593 | -84.1365
0.82963 | -84.1413
0.83333 | -84.417
0.83704 | -84.816
0.84074 | -84.8088
0.84444 | -84.3868
0.84815 | -84.2573
0.85185 | -84.5742
0.85556 | -85.1086
0.85926 | -85.2915
0.86296 | -84.8867
0.86667 | -84.353
0.87037 | -84.2775
0.87407 | -84.7818
0.87778 | -84.923
0.88148 | -84.3471
0.88518 | -84.1239
0.88889 | -84.5454
0.89259 | -85.1331
0.8963 | -85.7326
0.9 | -86.1467
0.9037 | -85.7357
0.90741 | -85.0629
0.91111 | -85.0991
0.91482 | -85.5857
0.91852 | -85.5318
0.92222 | -85.3672
0.92593 | -85.4431
0.92963 | -85.7935
0.93333 | -86.3552
0.93704 | -86.6566
0.94074 | -86.8803
0.94444 | -87.0168
0.94815 | -87.0039
0.95185 | -86.83
0.95556 | -86.2397
0.95926 | -86.0157
0.96296 | -86.6303
0.96667 | -86.9092
0.97037 | -86.6185
0.97407 | -86.6126
0.97778 | -86.966
0.98148 | -87.4889
0.98518 | -87.7514
0.98889 | -87.3051
0.99259 | -86.9476
0.9963 | -86.7238
1 | -86.0554
1.0037 | -86.0137
1.0074 | -86.4385
1.0111 | -86.7219
1.0148 | -87.0423
1.0185 | -87.2036
1.0222 | -86.9345
1.0259 | -86.5002
1.0296 | -86.1942
1.0333 | -86.1674
1.037 | -86.5211
1.0407 | -87.0359
1.0444 | -86.7964
1.0481 | -86.2857
1.0519 | -86.5269
1.0556 | -87.0796
1.0593 | -86.9914
1.063 | -86.6674
1.0667 | -86.8949
1.0704 | -87.0233
1.0741 | -87.114
1.0778 | -87.2519
1.0815 | -87.452
1.0852 | -87.5255
1.0889 | -87.359
1.0926 | -87.0791
1.0963 | -86.8945
1.1 | -87.0126
1.1037 | -87.4039
1.1074 | -87.6985
1.1111 | -87.6397
1.1148 | -87.3494
1.1185 | -87.0991
1.1222 | -87.1023
1.1259 | -87.4202
1.1296 | -87.8886
1.1333 | -88.2014
1.137 | -88.2824
1.1407 | -88.2582
1.1444 | -88.238
1.1481 | -88.217
1.1519 | -88.0633
1.1556 | -87.7885
1.1593 | -87.6332
1.163 | -87.7566
1.1667 | -88.0355
1.1704 | -88.178
1.1741 | -88.1163
1.1778 | -87.9832
1.1815 | -87.8012
1.1852 | -87.5665
1.1889 | -87.3855
1.1926 | -87.3692
1.1963 | -87.5651
1.2 | -87.902
1.2037 | -88.1357
1.2074 | -88.1024
1.2111 | -88.0537
1.2148 | -88.2587
1.2185 | -88.6075
1.2222 | -88.8248
1.2259 | -88.7818
1.2296 | -88.5136
1.2333 | -88.1188
1.237 | -87.7271
1.2407 | -87.5164
1.2444 | -87.6024
1.2481 | -87.9016
1.2519 | -88.2948
1.2556 | -88.8132
1.2593 | -89.2651
1.263 | -89.2799
1.2667 | -89.0018
1.2704 | -88.7984
1.2741 | -88.7255
1.2778 | -88.7096
1.2815 | -88.7798
1.2852 | -88.9011
1.2889 | -88.8386
1.2926 | -88.6199
1.2963 | -88.5099
1.3 | -88.5883
1.3037 | -88.7685
1.3074 | -89.0468
1.3111 | -89.4835
1.3148 | -89.8706
1.3185 | -89.7994
1.3222 | -89.291
1.3259 | -88.812
1.3296 | -88.7079
1.3333 | -88.8627
1.337 | -88.9835
1.3407 | -89.0012
1.3444 | -88.9703
1.3481 | -88.8867
1.3519 | -88.782
1.3556 | -88.8102
1.3593 | -89.0131
1.363 | -89.2163
1.3667 | -89.2784
1.3704 | -89.3013
1.3741 | -89.4665
1.3778 | -89.6747
1.3815 | -89.6772
1.3852 | -89.5618
1.3889 | -89.6804
1.3926 | -90.1209
1.3963 | -90.4828
1.4 | -90.4658
1.4037 | -90.3809
1.4074 | -90.4247
1.4111 | -90.423
1.4148 | -90.2972
1.4185 | -90.0937
1.4222 | -89.796
1.4259 | -89.5434
1.4296 | -89.5398
1.4333 | -89.7148
1.437 | -89.8558
1.4407 | -89.9267
1.4444 | -89.977
1.4481 | -89.9418
1.4519 | -89.8175
1.4556 | -89.7908
1.4593 | -89.9653
1.463 | -90.1995
1.4667 | -90.275
1.4704 | -90.1539
1.4741 | -89.9901
1.4778 | -89.8886
1.4815 | -89.9509
1.4852 | -90.2926
1.4889 | -90.7301
1.4926 | -90.9059
1.4963 | -90.7086
1.5 | -90.264
1.5037 | -89.7734
1.5074 | -89.4907
1.5111 | -89.4958
1.5148 | -89.5533
1.5185 | -89.4835
1.5222 | -89.4722
1.5259 | -89.6693
1.5296 | -90.039
1.5333 | -90.4831
1.537 | -90.8892
1.5407 | -91.1426
1.5444 | -91.163
1.5481 | -91.0478
1.5519 | -90.9615
1.5556 | -90.8693
1.5593 | -90.7304
1.563 | -90.7212
1.5667 | -90.9216
1.5704 | -91.2103
1.5741 | -91.4562
1.5778 | -91.5599
1.5815 | -91.4283
1.5852 | -91.154
1.5889 | -90.8978
1.5926 | -90.6502
1.5963 | -90.4455
1.6 | -90.5803
1.6037 | -90.7174
1.6074 | -90.8571
1.6111 | -90.9109
1.6148 | -90.9116
1.6185 | -90.9654
1.6222 | -91.1253
1.6259 | -91.2959
1.6296 | -91.3477
1.6333 | -91.2517
1.637 | -91.1164
1.6407 | -91.0568
1.6444 | -91.1032
1.6481 | -91.185
1.6519 | -91.2282
1.6556 | -91.2729
1.6593 | -91.4012
1.663 | -91.5721
1.6667 | -91.6941
1.6704 | -91.7101
1.6741 | -91.6303
1.6778 | -91.5379
1.6815 | -91.5124
1.6852 | -91.5416
1.6889 | -91.5682
1.6926 | -91.616
1.6963 | -91.7409
1.7 | -91.8404
1.7037 | -91.7585
1.7074 | -91.4694
1.7111 | -91.1051
1.7148 | -90.8296
1.7185 | -90.737
1.7222 | -90.8438
1.7259 | -91.0725
1.7296 | -91.2747
1.7333 | -91.3213
1.737 | -91.1733
1.7407 | -90.9307
1.7444 | -90.7286
1.7481 | -90.6808
1.7519 | -90.8443
1.7556 | -91.1724
1.7593 | -91.5287
1.763 | -91.7536
1.7667 | -91.8111
1.7704 | -91.7949
1.7741 | -91.8187
1.7778 | -91.9537
1.7815 | -92.1542
1.7852 | -92.2891
1.7889 | -92.2629
1.7926 | -92.1298
1.7963 | -92.0301
1.8 | -92.038
1.8037 | -92.1456
1.8074 | -92.2742
1.8111 | -92.3458
1.8148 | -92.3219
1.8185 | -92.1649
1.8222 | -91.8745
1.8259 | -91.5384
1.8296 | -91.3052
1.8333 | -91.2607
1.837 | -91.3672
1.8407 | -91.4974
1.8444 | -91.5481
1.8481 | -91.5604
1.8519 | -91.678
1.8556 | -91.9254
1.8593 | -92.1669
1.863 | -92.2529
1.8667 | -92.2156
1.8704 | -92.1838
1.8741 | -92.2192
1.8778 | -92.3082
1.8815 | -92.3976
1.8852 | -92.445
1.8889 | -92.4517
1.8926 | -92.4193
1.8963 | -92.3839
1.9 | -92.4074
1.9037 | -92.526
1.9074 | -92.7325
1.9111 | -92.9378
1.9148 | -92.9982
1.9185 | -92.817
1.9222 | -92.4959
1.9259 | -92.2246
1.9296 | -92.0976
1.9333 | -92.1084
1.937 | -92.2428
1.9407 | -92.4852
1.9444 | -92.7511
1.9481 | -92.8641
1.9519 | -92.6845
1.9556 | -92.3027
1.9593 | -91.931
1.963 | -91.7426
1.9667 | -91.8034
1.9704 | -92.0772
1.9741 | -92.4223
1.9778 | -92.6571
1.9815 | -92.6816
1.9852 | -92.5772
1.9889 | -92.4808
1.9926 | -92.4885
1.9963 | -92.6082
2 | -92.7805
2.0037 | -92.9037
2.0074 | -92.9078
2.0111 | -92.8144
2.0148 | -92.7071
2.0185 | -92.61
2.0222 | -92.4962
2.0259 | -92.3997
2.0296 | -92.4447
2.0333 | -92.687
2.037 | -93.0351
2.0407 | -93.293
2.0444 | -93.35
2.0481 | -93.283
2.0519 | -93.2066
2.0556 | -93.144
2.0593 | -93.0899
2.063 | -93.0441
2.0667 | -92.9944
2.0704 | -92.9409
2.0741 | -92.9007
2.0778 | -92.8492
2.0815 | -92.7852
2.0852 | -92.7661
2.0889 | -92.8625
2.0926 | -93.0407
2.0963 | -93.2062
2.1 | -93.2762
2.1037 | -93.2041
2.1074 | -93.0298
2.1111 | -92.8656
2.1148 | -92.8184
2.1185 | -92.9183
2.1222 | -93.0863
2.1259 | -93.2144
2.1296 | -93.2361
2.1333 | -93.1687
2.137 | -93.1255
2.1407 | -93.1309
2.1444 | -93.2023
2.1481 | -93.3311
2.1519 | -93.4897
2.1556 | -93.6415
2.1593 | -93.7526
2.163 | -93.8028
2.1667 | -93.7711
2.1704 | -93.6865
2.1741 | -93.5892
2.1778 | -93.4909
2.1815 | -93.3737
2.1852 | -93.2437
2.1889 | -93.1319
2.1926 | -93.0736
2.1963 | -93.0807
2.2 | -93.1215
2.2037 | -93.1873
2.2074 | -93.3048
2.2111 | -93.4714
2.2148 | -93.6643
2.2185 | -93.8032
2.2222 | -93.8209
2.2259 | -93.7199
2.2296 | -93.5556
2.2333 | -93.4094
2.237 | -93.3075
2.2407 | -93.2575
2.2444 | -93.2564
2.2481 | -93.3028
2.2519 | -93.3903
2.2556 | -93.4865
2.2593 | -93.5795
2.263 | -93.6886
2.2667 | -93.8162
2.2704 | -93.9686
2.2741 | -94.101
2.2778 | -94.1584
2.2815 | -94.112
2.2852 | -93.9847
2.2889 | -93.8355
2.2926 | -93.7229
2.2963 | -93.6749
2.3 | -93.6724
2.3037 | -93.6686
2.3074 | -93.6542
2.3111 | -93.6319
2.3148 | -93.6117
2.3185 | -93.6249
2.3222 | -93.6549
2.3259 | -93.6814
2.3296 | -93.6928
2.3333 | -93.6931
2.337 | -93.7341
2.3407 | -93.8578
2.3444 | -94.0434
2.3481 | -94.1952
2.3519 | -94.2214
2.3556 | -94.1546
2.3593 | -94.073
2.363 | -94.0654
2.3667 | -94.1565
2.3704 | -94.3155
2.3741 | -94.4932
2.3778 | -94.6155
2.3815 | -94.6292
2.3852 | -94.5549
2.3889 | -94.4381
2.3926 | -94.3332
2.3963 | -94.2703
2.4 | -94.27
2.4037 | -94.335
2.4074 | -94.4438
2.4111 | -94.5705
2.4148 | -94.6522
2.4185 | -94.6508
2.4222 | -94.5948
2.4259 | -94.5338
2.4296 | -94.5111
2.4333 | -94.529
2.437 | -94.5648
2.4407 | -94.5871
2.4444 | -94.5702
2.4481 | -94.5282
2.4519 | -94.4907
2.4556 | -94.4936
2.4593 | -94.5526
2.463 | -94.6429
2.4667 | -94.73
2.4704 | -94.7888
2.4741 | -94.8041
2.4778 | -94.7727
2.4815 | -94.7248
2.4852 | -94.672
2.4889 | -94.6242
2.4926 | -94.5768
2.4963 | -94.5075
2.5 | -94.4145
2.5037 | -94.3346
2.5074 | -94.2937
2.5111 | -94.3112
2.5148 | -94.354
2.5185 | -94.3911
2.5222 | -94.3955
2.5259 | -94.3865
2.5296 | -94.4034
2.5333 | -94.4817
2.537 | -94.5905
2.5407 | -94.6519
2.5444 | -94.6134
2.5481 | -94.4827
2.5519 | -94.2976
2.5556 | -94.1434
2.5593 | -94.0628
2.563 | -94.0837
2.5667 | -94.1938
2.5704 | -94.3482
2.5741 | -94.5009
2.5778 | -94.6515
2.5815 | -94.8025
2.5852 | -94.9364
2.5889 | -95.0231
2.5926 | -95.0503
2.5963 | -95.0231
2.6 | -95.0009
2.6037 | -95.0123
2.6074 | -95.0747
2.6111 | -95.1952
2.6148 | -95.3502
2.6185 | -95.4966
2.6222 | -95.5825
2.6259 | -95.5717
2.6296 | -95.4825
2.6333 | -95.371
2.637 | -95.2751
2.6407 | -95.2045
2.6444 | -95.1308
2.6481 | -95.0176
2.6519 | -94.8637
2.6556 | -94.6883
2.6593 | -94.5238
2.663 | -94.4007
2.6667 | -94.3294
2.6704 | -94.3827
2.6741 | -94.4266
2.6778 | -94.5178
2.6815 | -94.6327
2.6852 | -94.7347
2.6889 | -94.812
2.6926 | -94.8685
2.6963 | -94.9149
2.7 | -94.9693
2.7037 | -95.0377
2.7074 | -95.1049
2.7111 | -95.1626
2.7148 | -95.1922
2.7185 | -95.1925
2.7222 | -95.1547
2.7259 | -95.0859
2.7296 | -95.009
2.7333 | -94.929
2.737 | -94.8634
2.7407 | -94.8236
2.7444 | -94.8183
2.7481 | -94.8463
2.7519 | -94.8944
2.7556 | -94.959
2.7593 | -95.017
2.763 | -95.0465
2.7667 | -95.0615
2.7704 | -95.0878
2.7741 | -95.1594
2.7778 | -95.2751
2.7815 | -95.4189
2.7852 | -95.542
2.7889 | -95.6024
2.7926 | -95.5778
2.7963 | -95.4756
2.8 | -95.3337
2.8037 | -95.1961
2.8074 | -95.116
2.8111 | -95.1245
2.8148 | -95.2189
2.8185 | -95.3731
2.8222 | -95.5482
2.8259 | -95.6845
2.8296 | -95.7286
2.8333 | -95.6794
2.837 | -95.5746
2.8407 | -95.4654
2.8444 | -95.3946
2.8481 | -95.371
2.8519 | -95.3789
2.8556 | -95.392
2.8593 | -95.41
2.863 | -95.4358
2.8667 | -95.4702
2.8704 | -95.5007
2.8741 | -95.5154
2.8778 | -95.5006
2.8815 | -95.4553
2.8852 | -95.393
2.8889 | -95.341
2.8926 | -95.3185
2.8963 | -95.3383
2.9 | -95.3911
2.9037 | -95.4567
2.9074 | -95.4978
2.9111 | -95.5025
2.9148 | -95.4861
2.9185 | -95.4566
2.9222 | -95.4289
2.9259 | -95.3998
2.9296 | -95.3785
2.9333 | -95.364
2.937 | -95.3666
2.9407 | -95.3793
2.9444 | -95.3936
2.9481 | -95.414
2.9519 | -95.4568
2.9556 | -95.5349
2.9593 | -95.6385
2.963 | -95.7531
2.9667 | -95.8416
2.9704 | -95.8637
2.9741 | -95.8178
2.9778 | -95.7292
2.9815 | -95.6548
2.9852 | -95.606
2.9889 | -95.5905
2.9926 | -95.6087
2.9963 | -95.6543
3 | -95.7155
3.0037 | -95.7899
3.0074 | -95.855
3.0111 | -95.8638
3.0148 | -95.8157
3.0185 | -95.734
3.0222 | -95.6623
3.0259 | -95.6392
3.0296 | -95.6846
3.0333 | -95.7993
3.037 | -95.9407
3.0407 | -96.066
3.0444 | -96.137
3.0481 | -96.1474
3.0519 | -96.1144
3.0556 | -96.0599
3.0593 | -96.0233
3.063 | -96.0195
3.0667 | -96.0539
3.0704 | -96.115
3.0741 | -96.1867
3.0778 | -96.2504
3.0815 | -96.3083
3.0852 | -96.3647
3.0889 | -96.4113
3.0926 | -96.4564
3.0963 | -96.5093
3.1 | -96.5665
3.1037 | -96.6406
3.1074 | -96.7183
3.1111 | -96.7923
3.1148 | -96.8424
3.1185 | -96.852
3.1222 | -96.8234
3.1259 | -96.7735
3.1296 | -96.6939
3.1333 | -96.5812
3.137 | -96.4441
3.1407 | -96.3044
3.1444 | -96.1953
3.1481 | -96.1632
3.1519 | -96.2143
3.1556 | -96.354
3.1593 | -96.532
3.163 | -96.6745
3.1667 | -96.7222
3.1704 | -96.645
3.1741 | -96.4834
3.1778 | -96.3182
3.1815 | -96.2103
3.1852 | -96.1674
3.1889 | -96.1886
3.1926 | -96.2492
3.1963 | -96.309
3.2 | -96.3341
3.2037 | -96.2973
3.2074 | -96.2918
3.2111 | -96.2953
3.2148 | -96.3186
3.2185 | -96.3684
3.2222 | -96.4545
3.2259 | -96.5698
3.2296 | -96.6869
3.2333 | -96.7785
3.237 | -96.8395
3.2407 | -96.8463
3.2444 | -96.8064
3.2481 | -96.7552
3.2519 | -96.741
3.2556 | -96.7655
3.2593 | -96.838
3.263 | -96.9623
3.2667 | -97.1082
3.2704 | -97.2475
3.2741 | -97.3483
3.2778 | -97.3917
3.2815 | -97.3812
3.2852 | -97.3359
3.2889 | -97.2564
3.2926 | -97.1595
3.2963 | -97.0677
3.3 | -96.9883
3.3037 | -96.9222
3.3074 | -96.8603
3.3111 | -96.8263
3.3148 | -96.8141
3.3185 | -96.8193
3.3222 | -96.8507
3.3259 | -96.9061
3.3296 | -96.9784
3.3333 | -97.0529
3.337 | -97.1078
3.3407 | -97.1419
3.3444 | -97.1669
3.3481 | -97.186
3.3519 | -97.2224
3.3556 | -97.2753
3.3593 | -97.335
3.363 | -97.3887
3.3667 | -97.4064
3.3704 | -97.3783
3.3741 | -97.3023
3.3778 | -97.1891
3.3815 | -97.0603
3.3852 | -96.9483
3.3889 | -96.8779
3.3926 | -96.8471
3.3963 | -96.8353
3.4 | -96.8126
3.4037 | -96.7667
3.4074 | -96.6867
3.4111 | -96.5726
3.4148 | -96.4571
3.4185 | -96.3526
3.4222 | -96.2803
3.4259 | -96.2457
3.4296 | -96.2436
3.4333 | -96.2594
3.437 | -96.2889
3.4407 | -96.3324
3.4444 | -96.4017
3.4481 | -96.4985
3.4519 | -96.6127
3.4556 | -96.7282
3.4593 | -96.8185
3.463 | -96.872
3.4667 | -96.9045
3.4704 | -96.9144
3.4741 | -96.9176
3.4778 | -96.913
3.4815 | -96.9017
3.4852 | -96.8902
3.4889 | -96.8865
3.4926 | -96.9067
3.4963 | -96.9637
3.5 | -97.0558
3.5037 | -97.1823
3.5074 | -97.3154
3.5111 | -97.4277
3.5148 | -97.4988
3.5185 | -97.5071
3.5222 | -97.4665
3.5259 | -97.3953
3.5296 | -97.3426
3.5333 | -97.3226
3.537 | -97.3396
3.5407 | -97.3939
3.5444 | -97.478
3.5481 | -97.5768
3.5519 | -97.6776
3.5556 | -97.7732
3.5593 | -97.8438
3.563 | -97.8882
3.5667 | -97.9119
3.5704 | -97.9253
3.5741 | -97.9482
3.5778 | -97.9698
3.5815 | -97.9837
3.5852 | -97.9663
3.5889 | -97.9081
3.5926 | -97.8021
3.5963 | -97.652
3.6 | -97.4817
3.6037 | -97.3254
3.6074 | -97.2062
3.6111 | -97.1292
3.6148 | -97.0954
3.6185 | -97.1035
3.6222 | -97.1366
3.6259 | -97.1855
3.6296 | -97.2474
3.6333 | -97.31
3.637 | -97.3616
3.6407 | -97.4072
3.6444 | -97.4426
3.6481 | -97.4613
3.6519 | -97.4678
3.6556 | -97.4663
3.6593 | -97.4703
3.663 | -97.481
3.6667 | -97.5031
3.6704 | -97.5281
3.6741 | -97.5488
3.6778 | -97.5748
3.6815 | -97.6179
3.6852 | -97.6759
3.6889 | -97.75
3.6926 | -97.8292
3.6963 | -97.8844
3.7 | -97.9101
3.7037 | -97.9111
3.7074 | -97.9222
3.7111 | -97.945
3.7148 | -97.9787
3.7185 | -98.0384
3.7222 | -98.0836
3.7259 | -98.1031
3.7296 | -98.0947
3.7333 | -98.0737
3.737 | -98.0548
3.7407 | -98.0496
3.7444 | -98.0566
3.7481 | -98.082
3.7519 | -98.1204
3.7556 | -98.1746
3.7593 | -98.2344
3.763 | -98.2855
3.7667 | -98.3056
3.7704 | -98.2986
3.7741 | -98.2658
3.7778 | -98.222
3.7815 | -98.1842
3.7852 | -98.145
3.7889 | -98.0997
3.7926 | -98.0569
3.7963 | -98.0122
3.8 | -97.9726
3.8037 | -97.943
3.8074 | -97.9234
3.8111 | -97.9096
3.8148 | -97.9017
3.8185 | -97.895
3.8222 | -97.8883
3.8259 | -97.8679
3.8296 | -97.8172
3.8333 | -97.747
3.837 | -97.6672
3.8407 | -97.587
3.8444 | -97.5149
3.8481 | -97.4626
3.8519 | -97.4355
3.8556 | -97.4192
3.8593 | -97.4179
3.863 | -97.4221
3.8667 | -97.4335
3.8704 | -97.4566
3.8741 | -97.5043
3.8778 | -97.5781
3.8815 | -97.678
3.8852 | -97.7963
3.8889 | -97.9352
3.8926 | -98.0943
3.8963 | -98.2511
3.9 | -98.3802
3.9037 | -98.4769
3.9074 | -98.5409
3.9111 | -98.5523
3.9148 | -98.5122
3.9185 | -98.4399
3.9222 | -98.3462
3.9259 | -98.244
3.9296 | -98.1452
3.9333 | -98.0698
3.937 | -98.0223
3.9407 | -98.0167
3.9444 | -98.0569
3.9481 | -98.1323
3.9519 | -98.2276
3.9556 | -98.3289
3.9593 | -98.4315
3.963 | -98.5244
3.9667 | -98.5953
3.9704 | -98.6389
3.9741 | -98.6515
3.9778 | -98.6342
3.9815 | -98.6024
3.9852 | -98.5719
3.9889 | -98.5543
3.9926 | -98.5476
3.9963 | -98.5493
4 | -98.556
4.0037 | -98.5467
4.0074 | -98.5125
4.0111 | -98.4698
4.0148 | -98.4139
4.0185 | -98.3558
4.0222 | -98.2994
4.0259 | -98.2481
4.0296 | -98.2016
4.0333 | -98.1674
4.037 | -98.1481
4.0407 | -98.1389
4.0444 | -98.1286
4.0481 | -98.1161
4.0519 | -98.0963
4.0556 | -98.0761
4.0593 | -98.0537
4.063 | -98.0344
4.0667 | -98.0353
4.0704 | -98.0603
4.0741 | -98.117
4.0778 | -98.2004
4.0815 | -98.2917
4.0852 | -98.3899
4.0889 | -98.4839
4.0926 | -98.5554
4.0963 | -98.5898
4.1 | -98.5991
4.1037 | -98.5926
4.1074 | -98.5749
4.1111 | -98.5395
4.1148 | -98.5002
4.1185 | -98.4666
4.1222 | -98.4443
4.1259 | -98.4324
4.1296 | -98.4145
4.1333 | -98.3927
4.137 | -98.3501
4.1407 | -98.2848
4.1444 | -98.2064
4.1481 | -98.1198
4.1519 | -98.0492
4.1556 | -98.0213
4.1593 | -98.0384
4.163 | -98.091
4.1667 | -98.1803
4.1704 | -98.2964
4.1741 | -98.4235
4.1778 | -98.5481
4.1815 | -98.6463
4.1852 | -98.7142
4.1889 | -98.7503
4.1926 | -98.748
4.1963 | -98.6983
4.2 | -98.621
4.2037 | -98.536
4.2074 | -98.4434
4.2111 | -98.3468
4.2148 | -98.2557
4.2185 | -98.1714
4.2222 | -98.1168
4.2259 | -98.1079
4.2296 | -98.1335
4.2333 | -98.1936
4.237 | -98.2882
4.2407 | -98.3977
4.2444 | -98.5118
4.2481 | -98.6126
4.2519 | -98.6827
4.2556 | -98.7177
4.2593 | -98.7109
4.263 | -98.6647
4.2667 | -98.5874
4.2704 | -98.4943
4.2741 | -98.4236
4.2778 | -98.3734
4.2815 | -98.3437
4.2852 | -98.348
4.2889 | -98.3862
4.2926 | -98.4478
4.2963 | -98.5372
4.3 | -98.6458
4.3037 | -98.7464
4.3074 | -98.823
4.3111 | -98.8733
4.3148 | -98.8951
4.3185 | -98.894
4.3222 | -98.8791
4.3259 | -98.8679
4.3296 | -98.8693
4.3333 | -98.8848
4.337 | -98.9105
4.3407 | -98.9393
4.3444 | -98.9653
4.3481 | -98.988
4.3519 | -99.0045
4.3556 | -99.0149
4.3593 | -99.016
4.363 | -99.0234
4.3667 | -99.0325
4.3704 | -99.0342
4.3741 | -99.0419
4.3778 | -99.0576
4.3815 | -99.0721
4.3852 | -99.0913
4.3889 | -99.1214
4.3926 | -99.1652
4.3963 | -99.2105
4.4 | -99.2558
4.4037 | -99.303
4.4074 | -99.3501
4.4111 | -99.3913
4.4148 | -99.4196
4.4185 | -99.4258
4.4222 | -99.4194
4.4259 | -99.3917
4.4296 | -99.3533
4.4333 | -99.3071
4.437 | -99.262
4.4407 | -99.2292
4.4444 | -99.2147
4.4481 | -99.2117
4.4519 | -99.2222
4.4556 | -99.2347
4.4593 | -99.2456
4.463 | -99.2442
4.4667 | -99.2389
4.4704 | -99.2292
4.4741 | -99.2077
4.4778 | -99.1746
4.4815 | -99.1383
4.4852 | -99.1018
4.4889 | -99.0774
4.4926 | -99.0674
4.4963 | -99.0744
4.5 | -99.102
4.5037 | -99.1467
4.5074 | -99.2072
4.5111 | -99.2701
4.5148 | -99.322
4.5185 | -99.3699
4.5222 | -99.3956
4.5259 | -99.3975
4.5296 | -99.3723
4.5333 | -99.3288
4.537 | -99.2629
4.5407 | -99.1905
4.5444 | -99.1302
4.5481 | -99.086
4.5519 | -99.0751
4.5556 | -99.105
4.5593 | -99.1656
4.563 | -99.2563
4.5667 | -99.3676
4.5704 | -99.4872
4.5741 | -99.6087
4.5778 | -99.7145
4.5815 | -99.7931
4.5852 | -99.8539
4.5889 | -99.8832
4.5926 | -99.8854
4.5963 | -99.8531
4.6 | -99.7836
4.6037 | -99.6927
4.6074 | -99.5836
4.6111 | -99.4639
4.6148 | -99.3502
4.6185 | -99.2474
4.6222 | -99.1505
4.6259 | -99.0788
4.6296 | -99.0306
4.6333 | -98.99
4.637 | -98.9598
4.6407 | -98.9325
4.6444 | -98.907
4.6481 | -98.8835
4.6519 | -98.8704
4.6556 | -98.857
4.6593 | -98.8535
4.663 | -98.8656
4.6667 | -98.8941
4.6704 | -98.938
4.6741 | -98.992
4.6778 | -99.0602
4.6815 | -99.1421
4.6852 | -99.2379
4.6889 | -99.343
4.6926 | -99.4516
4.6963 | -99.5449
4.7 | -99.6243
4.7037 | -99.6859
4.7074 | -99.7135
4.7111 | -99.7097
4.7148 | -99.6851
4.7185 | -99.6526
4.7222 | -99.6163
4.7259 | -99.5867
4.7296 | -99.5687
4.7333 | -99.5602
4.737 | -99.5609
4.7407 | -99.5731
4.7444 | -99.5957
4.7481 | -99.6297
4.7519 | -99.6744
4.7556 | -99.7205
4.7593 | -99.7739
4.763 | -99.8176
4.7667 | -99.8485
4.7704 | -99.8617
4.7741 | -99.8557
4.7778 | -99.8375
4.7815 | -99.8195
4.7852 | -99.7954
4.7889 | -99.7718
4.7926 | -99.7528
4.7963 | -99.745
4.8 | -99.7413
4.8037 | -99.7444
4.8074 | -99.7727
4.8111 | -99.7957
4.8148 | -99.8232
4.8185 | -99.8604
4.8222 | -99.9069
4.8259 | -99.9522
4.8296 | -100.0049
4.8333 | -100.0574
4.837 | -100.1016
4.8407 | -100.1385
4.8444 | -100.1678
4.8481 | -100.1844
4.8519 | -100.1799
4.8556 | -100.1462
4.8593 | -100.0904
4.863 | -100.0181
4.8667 | -99.9305
4.8704 | -99.8396
4.8741 | -99.7631
4.8778 | -99.6961
4.8815 | -99.65
4.8852 | -99.626
4.8889 | -99.6191
4.8926 | -99.628
4.8963 | -99.643
4.9 | -99.6599
4.9037 | -99.6808
4.9074 | -99.699
4.9111 | -99.7228
4.9148 | -99.7551
4.9185 | -99.7887
4.9222 | -99.8177
4.9259 | -99.8537
4.9296 | -99.8896
4.9333 | -99.9237
4.937 | -99.9521
4.9407 | -99.9656
4.9444 | -99.9688
4.9481 | -99.9569
4.9519 | -99.9349
4.9556 | -99.9164
4.9593 | -99.8928
4.963 | -99.871
4.9667 | -99.8625
4.9704 | -99.8584
4.9741 | -99.8543
4.9778 | -99.8599
4.9815 | -99.8718
4.9852 | -99.8786
4.9889 | -99.8894
4.9926 | -99.9101
4.9963 | -99.9366
5 | -99.9755
5.0037 | -100.0119
5.0074 | -100.0487
5.0111 | -100.0953
5.0148 | -100.1443
5.0185 | -100.1931
5.0222 | -100.2369
5.0259 | -100.2762
5.0296 | -100.307
5.0333 | -100.3323
5.037 | -100.3484
5.0407 | -100.3529
5.0444 | -100.3337
5.0481 | -100.3003
5.0519 | -100.2601
5.0556 | -100.2043
5.0593 | -100.1412
5.063 | -100.0902
5.0667 | -100.0378
5.0704 | -99.9948
5.0741 | -99.9681
5.0778 | -99.9575
5.0815 | -99.9676
5.0852 | -99.9925
5.0889 | -100.0252
5.0926 | -100.0795
5.0963 | -100.1417
5.1 | -100.2109
5.1037 | -100.2774
5.1074 | -100.3269
5.1111 | -100.3515
5.1148 | -100.3689
5.1185 | -100.3685
5.1222 | -100.3549
5.1259 | -100.3202
5.1296 | -100.276
5.1333 | -100.2238
5.137 | -100.172
5.1407 | -100.1292
5.1444 | -100.1054
5.1481 | -100.1075
5.1519 | -100.1361
5.1556 | -100.1883
5.1593 | -100.2586
5.163 | -100.3489
5.1667 | -100.4446
5.1704 | -100.5368
5.1741 | -100.6279
5.1778 | -100.7014
5.1815 | -100.7631
5.1852 | -100.8067
5.1889 | -100.8221
5.1926 | -100.8001
5.1963 | -100.762
5.2 | -100.7155
5.2037 | -100.6694
5.2074 | -100.621
5.2111 | -100.5861
5.2148 | -100.5716
5.2185 | -100.569
5.2222 | -100.5739
5.2259 | -100.5841
5.2296 | -100.6069
5.2333 | -100.6398
5.237 | -100.6786
5.2407 | -100.7167
5.2444 | -100.7541
5.2481 | -100.7895
5.2519 | -100.8173
5.2556 | -100.8396
5.2593 | -100.8492
5.263 | -100.8662
5.2667 | -100.8845
5.2704 | -100.8979
5.2741 | -100.9076
5.2778 | -100.9122
5.2815 | -100.9165
5.2852 | -100.9203
5.2889 | -100.915
5.2926 | -100.8971
5.2963 | -100.8728
5.3 | -100.8521
5.3037 | -100.8116
5.3074 | -100.7428
5.3111 | -100.6573
5.3148 | -100.5583
5.3185 | -100.4482
5.3222 | -100.3286
5.3259 | -100.2097
5.3296 | -100.1081
5.3333 | -100.0274
5.337 | -99.9726
5.3407 | -100.0047
5.3444 | -100.0024
5.3481 | -100.0245
5.3519 | -100.0696
5.3556 | -100.1295
5.3593 | -100.1982
5.363 | -100.2666
5.3667 | -100.3324
5.3704 | -100.398
5.3741 | -100.4616
5.3778 | -100.5226
5.3815 | -100.5787
5.3852 | -100.6407
5.3889 | -100.7104
5.3926 | -100.7763
5.3963 | -100.8355
5.4 | -100.8812
5.4037 | -100.9297
5.4074 | -100.9773
5.4111 | -101.0059
5.4148 | -101.0228
5.4185 | -101.0235
5.4222 | -101.0041
5.4259 | -100.9691
5.4296 | -100.9254
5.4333 | -100.8674
5.437 | -100.801
5.4407 | -100.7322
5.4444 | -100.6731
5.4481 | -100.6217
5.4519 | -100.5766
5.4556 | -100.544
5.4593 | -100.5175
5.463 | -100.4916
5.4667 | -100.4704
5.4704 | -100.4499
5.4741 | -100.4221
5.4778 | -100.3944
5.4815 | -100.3762
5.4852 | -100.3646
5.4889 | -100.3512
5.4926 | -100.3442
5.4963 | -100.3523
5.5 | -100.3785
5.5037 | -100.4245
5.5074 | -100.477
5.5111 | -100.5326
5.5148 | -100.5915
5.5185 | -100.6467
5.5222 | -100.696
5.5259 | -100.733
5.5296 | -100.7607
5.5333 | -100.7809
5.537 | -100.8002
5.5407 | -100.8122
5.5444 | -100.8245
5.5481 | -100.8451
5.5519 | -100.8638
5.5556 | -100.8857
5.5593 | -100.9152
5.563 | -100.9468
5.5667 | -100.9689
5.5704 | -100.9857
5.5741 | -101.01
5.5778 | -101.035
5.5815 | -101.0488
5.5852 | -101.0484
5.5889 | -101.0381
5.5926 | -101.0266
5.5963 | -101.003
5.6 | -100.9721
5.6037 | -100.946
5.6074 | -100.924
5.6111 | -100.911
5.6148 | -100.9096
5.6185 | -100.9274
5.6222 | -100.9547
5.6259 | -100.9869
5.6296 | -101.0122
5.6333 | -101.0282
5.637 | -101.0322
5.6407 | -101.021
5.6444 | -101.0004
5.6481 | -100.9733
5.6519 | -100.949
5.6556 | -100.9272
5.6593 | -100.8972
5.663 | -100.8755
5.6667 | -100.871
5.6704 | -100.869
5.6741 | -100.8674
5.6778 | -100.8683
5.6815 | -100.8689
5.6852 | -100.8646
5.6889 | -100.8515
5.6926 | -100.8338
5.6963 | -100.8185
5.7 | -100.8032
5.7037 | -100.7862
5.7074 | -100.7715
5.7111 | -100.7704
5.7148 | -100.7786
5.7185 | -100.7977
5.7222 | -100.8249
5.7259 | -100.8564
5.7296 | -100.8965
5.7333 | -100.943
5.737 | -100.9902
5.7407 | -101.0398
5.7444 | -101.0804
5.7481 | -101.107
5.7519 | -101.1192
5.7556 | -101.1255
5.7593 | -101.1217
5.763 | -101.1019
5.7667 | -101.0798
5.7704 | -101.0559
5.7741 | -101.0301
5.7778 | -101.0177
5.7815 | -101.0132
5.7852 | -101.018
5.7889 | -101.04
5.7926 | -101.0683
5.7963 | -101.0973
5.8 | -101.1321
5.8037 | -101.157
5.8074 | -101.1692
5.8111 | -101.1684
5.8148 | -101.152
5.8185 | -101.1254
5.8222 | -101.0937
5.8259 | -101.0614
5.8296 | -101.034
5.8333 | -101.0086
5.837 | -100.9931
5.8407 | -100.985
5.8444 | -100.9887
5.8481 | -101.0052
5.8519 | -101.0303
5.8556 | -101.0611
5.8593 | -101.1002
5.863 | -101.1401
5.8667 | -101.1905
5.8704 | -101.2468
5.8741 | -101.3083
5.8778 | -101.3658
5.8815 | -101.4303
5.8852 | -101.4875
5.8889 | -101.5354
5.8926 | -101.5793
5.8963 | -101.6234
5.9 | -101.6585
5.9037 | -101.6775
5.9074 | -101.6879
5.9111 | -101.6815
5.9148 | -101.6549
5.9185 | -101.6133
5.9222 | -101.5518
5.9259 | -101.473
5.9296 | -101.3853
5.9333 | -101.2865
5.937 | -101.1897
5.9407 | -101.0989
5.9444 | -101.0193
5.9481 | -100.955
5.9519 | -100.9121
5.9556 | -100.893
5.9593 | -100.8967
5.963 | -100.922
5.9667 | -100.9701
5.9704 | -101.0336
5.9741 | -101.1097
5.9778 | -101.198
5.9815 | -101.2915
5.9852 | -101.3772
5.9889 | -101.4511
5.9926 | -101.5149
5.9963 | -101.5574
6 | -101.572
6.0037 | -101.5655
6.0074 | -101.5329
6.0111 | -101.4808
6.0148 | -101.4196
6.0185 | -101.3496
6.0222 | -101.2753
6.0259 | -101.2078
6.0296 | -101.1372
6.0333 | -101.0721
6.037 | -101.0195
6.0407 | -100.9766
6.0444 | -100.9379
6.0481 | -100.9069
6.0519 | -100.8934
6.0556 | -100.8886
6.0593 | -100.8911
6.063 | -100.9083
6.0667 | -100.9314
6.0704 | -100.9592
6.0741 | -100.9974
6.0778 | -101.0515
6.0815 | -101.1081
6.0852 | -101.1776
6.0889 | -101.2504
6.0926 | -101.3245
6.0963 | -101.3924
6.1 | -101.4578
6.1037 | -101.5085
6.1074 | -101.535
6.1111 | -101.5409
6.1148 | -101.5296
6.1185 | -101.5084
6.1222 | -101.486
6.1259 | -101.4509
6.1296 | -101.411
6.1333 | -101.3749
6.137 | -101.3432
6.1407 | -101.3187
6.1444 | -101.2905
6.1481 | -101.2681
6.1519 | -101.2545
6.1556 | -101.2462
6.1593 | -101.2527
6.163 | -101.2739
6.1667 | -101.3
6.1704 | -101.3365
6.1741 | -101.3933
6.1778 | -101.4524
6.1815 | -101.5282
6.1852 | -101.6044
6.1889 | -101.6874
6.1926 | -101.7692
6.1963 | -101.8509
6.2 | -101.9288
6.2037 | -101.9942
6.2074 | -102.0514
6.2111 | -102.0928
6.2148 | -102.1197
6.2185 | -102.1388
6.2222 | -102.1474
6.2259 | -102.1463
6.2296 | -102.1379
6.2333 | -102.1197
6.237 | -102.0976
6.2407 | -102.0694
6.2444 | -102.0306
6.2481 | -101.9858
6.2519 | -101.9271
6.2556 | -101.8677
6.2593 | -101.8127
6.263 | -101.7454
6.2667 | -101.675
6.2704 | -101.6079
6.2741 | -101.5446
6.2778 | -101.4922
6.2815 | -101.4465
6.2852 | -101.4126
6.2889 | -101.3988
6.2926 | -101.4017
6.2963 | -101.4216
6.3 | -101.4487
6.3037 | -101.4798
6.3074 | -101.5156
6.3111 | -101.5552
6.3148 | -101.5905
6.3185 | -101.6139
6.3222 | -101.6384
6.3259 | -101.6616
6.3296 | -101.6775
6.3333 | -101.6866
6.337 | -101.6878
6.3407 | -101.6803
6.3444 | -101.6698
6.3481 | -101.6486
6.3519 | -101.6287
6.3556 | -101.6238
6.3593 | -101.622
6.363 | -101.6306
6.3667 | -101.6432
6.3704 | -101.661
6.3741 | -101.6838
6.3778 | -101.7091
6.3815 | -101.7335
6.3852 | -101.7543
6.3889 | -101.7611
6.3926 | -101.7567
6.3963 | -101.7411
6.4 | -101.7168
6.4037 | -101.685
6.4074 | -101.6506
6.4111 | -101.6116
6.4148 | -101.5847
6.4185 | -101.5536
6.4222 | -101.5262
6.4259 | -101.5107
6.4296 | -101.5061
6.4333 | -101.5142
6.437 | -101.5376
6.4407 | -101.5734
6.4444 | -101.617
6.4481 | -101.6635
6.4519 | -101.7175
6.4556 | -101.7764
6.4593 | -101.8344
6.463 | -101.8932
6.4667 | -101.952
6.4704 | -102.0011
6.4741 | -102.0388
6.4778 | -102.0685
6.4815 | -102.0971
6.4852 | -102.124
6.4889 | -102.1453
6.4926 | -102.1617
6.4963 | -102.1671
6.5 | -102.169
6.5037 | -102.176
6.5074 | -102.182
6.5111 | -102.1912
6.5148 | -102.1994
6.5185 | -102.1989
6.5222 | -102.1916
6.5259 | -102.1799
6.5296 | -102.1578
6.5333 | -102.1257
6.537 | -102.0842
6.5407 | -102.0407
6.5444 | -101.9938
6.5481 | -101.9456
6.5519 | -101.8931
6.5556 | -101.848
6.5593 | -101.8103
6.563 | -101.779
6.5667 | -101.7543
6.5704 | -101.7432
6.5741 | -101.7394
6.5778 | -101.7381
6.5815 | -101.7374
6.5852 | -101.7361
6.5889 | -101.7342
6.5926 | -101.7389
6.5963 | -101.7426
6.6 | -101.7439
6.6037 | -101.7457
6.6074 | -101.7512
6.6111 | -101.7576
6.6148 | -101.7713
6.6185 | -101.7928
6.6222 | -101.8189
6.6259 | -101.8449
6.6296 | -101.8786
6.6333 | -101.9195
6.637 | -101.9578
6.6407 | -101.9911
6.6444 | -102.0258
6.6481 | -102.0568
6.6519 | -102.0808
6.6556 | -102.0981
6.6593 | -102.1071
6.663 | -102.1128
6.6667 | -102.116
6.6704 | -102.1121
6.6741 | -102.1096
6.6778 | -102.1141
6.6815 | -102.1204
6.6852 | -102.1311
6.6889 | -102.1427
6.6926 | -102.1494
6.6963 | -102.1562
6.7 | -102.1598
6.7037 | -102.1595
6.7074 | -102.1645
6.7111 | -102.1651
6.7148 | -102.1592
6.7185 | -102.1545
6.7222 | -102.1468
6.7259 | -102.1361
6.7296 | -102.1222
6.7333 | -102.1091
6.737 | -102.104
6.7407 | -102.1109
6.7444 | -102.1219
6.7481 | -102.1403
6.7519 | -102.1699
6.7556 | -102.2033
6.7593 | -102.2453
6.763 | -102.2808
6.7667 | -102.305
6.7704 | -102.3261
6.7741 | -102.3276
6.7778 | -102.2981
6.7815 | -102.2591
6.7852 | -102.2076
6.7889 | -102.1437
6.7926 | -102.0758
6.7963 | -102.0044
6.8 | -101.9399
6.8037 | -101.8851
6.8074 | -101.8427
6.8111 | -101.816
6.8148 | -101.8085
6.8185 | -101.8193
6.8222 | -101.8484
6.8259 | -101.8934
6.8296 | -101.9495
6.8333 | -102.0186
6.837 | -102.0895
6.8407 | -102.1582
6.8444 | -102.2254
6.8481 | -102.2881
6.8519 | -102.3423
6.8556 | -102.3842
6.8593 | -102.4111
6.863 | -102.4275
6.8667 | -102.4239
6.8704 | -102.3987
6.8741 | -102.3677
6.8778 | -102.333
6.8815 | -102.2899
6.8852 | -102.244
6.8889 | -102.2012
6.8926 | -102.1665
6.8963 | -102.1437
6.9 | -102.1224
6.9037 | -102.1064
6.9074 | -102.0925
6.9111 | -102.0876
6.9148 | -102.0878
6.9185 | -102.088
6.9222 | -102.0908
6.9259 | -102.0998
6.9296 | -102.1115
6.9333 | -102.1218
6.937 | -102.1392
6.9407 | -102.1604
6.9444 | -102.191
6.9481 | -102.2181
6.9519 | -102.2505
6.9556 | -102.2865
6.9593 | -102.328
6.963 | -102.3708
6.9667 | -102.4174
6.9704 | -102.4704
6.9741 | -102.5237
6.9778 | -102.569
6.9815 | -102.6048
6.9852 | -102.6321
6.9889 | -102.659
6.9926 | -102.68
6.9963 | -102.6903
7 | -102.6919
7.0037 | -102.6936
7.0074 | -102.6897
7.0111 | -102.681
7.0148 | -102.6622
7.0185 | -102.6381
7.0222 | -102.6131
7.0259 | -102.5841
7.0296 | -102.5467
7.0333 | -102.5085
7.037 | -102.4704
7.0407 | -102.4414
7.0444 | -102.4163
7.0481 | -102.3993
7.0519 | -102.3905
7.0556 | -102.3903
7.0593 | -102.3991
7.063 | -102.4184
7.0667 | -102.448
7.0704 | -102.4851
7.0741 | -102.5302
7.0778 | -102.576
7.0815 | -102.6263
7.0852 | -102.6813
7.0889 | -102.7349
7.0926 | -102.7762
7.0963 | -102.8072
7.1 | -102.8275
7.1037 | -102.834
7.1074 | -102.8302
7.1111 | -102.8107
7.1148 | -102.7811
7.1185 | -102.748
7.1222 | -102.7057
7.1259 | -102.6542
7.1296 | -102.6012
7.1333 | -102.5511
7.137 | -102.5121
7.1407 | -102.4822
7.1444 | -102.4573
7.1481 | -102.4422
7.1519 | -102.4387
7.1556 | -102.4451
7.1593 | -102.4682
7.163 | -102.4993
7.1667 | -102.5329
7.1704 | -102.5754
7.1741 | -102.6145
7.1778 | -102.6478
7.1815 | -102.6737
7.1852 | -102.6887
7.1889 | -102.6849
7.1926 | -102.6643
7.1963 | -102.6339
7.2 | -102.5971
7.2037 | -102.551
7.2074 | -102.5109
7.2111 | -102.4725
7.2148 | -102.4376
7.2185 | -102.4064
7.2222 | -102.3812
7.2259 | -102.3639
7.2296 | -102.3583
7.2333 | -102.3593
7.237 | -102.364
7.2407 | -102.3774
7.2444 | -102.3949
7.2481 | -102.417
7.2519 | -102.4463
7.2556 | -102.4779
7.2593 | -102.5171
7.263 | -102.5589
7.2667 | -102.5956
7.2704 | -102.6292
7.2741 | -102.6632
7.2778 | -102.6872
7.2815 | -102.7148
7.2852 | -102.7395
7.2889 | -102.7575
7.2926 | -102.7687
7.2963 | -102.7821
7.3 | -102.7941
7.3037 | -102.7986
7.3074 | -102.7975
7.3111 | -102.792
7.3148 | -102.7799
7.3185 | -102.7692
7.3222 | -102.7573
7.3259 | -102.7455
7.3296 | -102.7309
7.3333 | -102.714
7.337 | -102.6932
7.3407 | -102.6702
7.3444 | -102.6404
7.3481 | -102.6065
7.3519 | -102.5768
7.3556 | -102.5487
7.3593 | -102.5213
7.363 | -102.496
7.3667 | -102.4791
7.3704 | -102.4724
7.3741 | -102.4705
7.3778 | -102.4764
7.3815 | -102.4871
7.3852 | -102.5043
7.3889 | -102.5205
7.3926 | -102.5382
7.3963 | -102.5545
7.4 | -102.5773
7.4037 | -102.5934
7.4074 | -102.6034
7.4111 | -102.6126
7.4148 | -102.623
7.4185 | -102.6317
7.4222 | -102.6346
7.4259 | -102.6358
7.4296 | -102.6418
7.4333 | -102.6503
7.437 | -102.6588
7.4407 | -102.6654
7.4444 | -102.6762
7.4481 | -102.6906
7.4519 | -102.7086
7.4556 | -102.7221
7.4593 | -102.7309
7.463 | -102.741
7.4667 | -102.7527
7.4704 | -102.7657
7.4741 | -102.7795
7.4778 | -102.8016
7.4815 | -102.8221
7.4852 | -102.8495
7.4889 | -102.8772
7.4926 | -102.9093
7.4963 | -102.9437
7.5 | -102.9778
7.5037 | -103.0048
7.5074 | -103.0291
7.5111 | -103.0466
7.5148 | -103.0619
7.5185 | -103.0776
7.5222 | -103.0875
7.5259 | -103.0869
7.5296 | -103.0766
7.5333 | -103.0576
7.537 | -103.0307
7.5407 | -102.993
7.5444 | -102.9475
7.5481 | -102.8985
7.5519 | -102.8453
7.5556 | -102.7882
7.5593 | -102.7308
7.563 | -102.6728
7.5667 | -102.6165
7.5704 | -102.5676
7.5741 | -102.523
7.5778 | -102.4872
7.5815 | -102.4576
7.5852 | -102.4357
7.5889 | -102.4247
7.5926 | -102.4153
7.5963 | -102.41
7.6 | -102.4138
7.6037 | -102.4274
7.6074 | -102.4406
7.6111 | -102.4554
7.6148 | -102.4707
7.6185 | -102.4883
7.6222 | -102.511
7.6259 | -102.5331
7.6296 | -102.5537
7.6333 | -102.5771
7.637 | -102.6016
7.6407 | -102.6302
7.6444 | -102.6689
7.6481 | -102.7137
7.6519 | -102.7573
7.6556 | -102.8006
7.6593 | -102.8468
7.663 | -102.8974
7.6667 | -102.9506
7.6704 | -103.0034
7.6741 | -103.0551
7.6778 | -103.1079
7.6815 | -103.1655
7.6852 | -103.2203
7.6889 | -103.2682
7.6926 | -103.3126
7.6963 | -103.3579
7.7 | -103.3996
7.7037 | -103.4302
7.7074 | -103.4583
7.7111 | -103.4758
7.7148 | -103.4827
7.7185 | -103.4773
7.7222 | -103.4645
7.7259 | -103.4464
7.7296 | -103.4185
7.7333 | -103.3843
7.737 | -103.3346
7.7407 | -103.2815
7.7444 | -103.2221
7.7481 | -103.1623
7.7519 | -103.1107
7.7556 | -103.065
7.7593 | -103.022
7.763 | -102.9891
7.7667 | -102.9703
7.7704 | -102.9604
7.7741 | -102.9609
7.7778 | -102.9716
7.7815 | -102.9902
7.7852 | -103.0192
7.7889 | -103.0515
7.7926 | -103.0872
7.7963 | -103.123
7.8 | -103.1505
7.8037 | -103.1788
7.8074 | -103.1972
7.8111 | -103.2051
7.8148 | -103.2062
7.8185 | -103.2045
7.8222 | -103.1977
7.8259 | -103.1935
7.8296 | -103.1826
7.8333 | -103.1747
7.837 | -103.167
7.8407 | -103.1581
7.8444 | -103.1527
7.8481 | -103.1508
7.8519 | -103.1523
7.8556 | -103.1585
7.8593 | -103.1695
7.863 | -103.1762
7.8667 | -103.1836
7.8704 | -103.1872
7.8741 | -103.1952
7.8778 | -103.2081
7.8815 | -103.2189
7.8852 | -103.2281
7.8889 | -103.2296
7.8926 | -103.2285
7.8963 | -103.2249
7.9 | -103.2121
7.9037 | -103.195
7.9074 | -103.1701
7.9111 | -103.1384
7.9148 | -103.1102
7.9185 | -103.0795
7.9222 | -103.0458
7.9259 | -103.0091
7.9296 | -102.9807
7.9333 | -102.9559
7.937 | -102.9307
7.9407 | -102.9085
7.9444 | -102.8926
7.9481 | -102.8808
7.9519 | -102.8794
7.9556 | -102.885
7.9593 | -102.8967
7.963 | -102.9105
7.9667 | -102.9318
7.9704 | -102.9582
7.9741 | -102.9946
7.9778 | -103.0384
7.9815 | -103.0835
7.9852 | -103.1321
7.9889 | -103.1775
7.9926 | -103.218
7.9963 | -103.2647
8 | -103.309
8.0037 | -103.3467
8.0074 | -103.3781
8.0111 | -103.4061
8.0148 | -103.418
8.0185 | -103.4355
8.0222 | -103.4535
8.0259 | -103.4706
8.0296 | -103.4813
8.0333 | -103.4895
8.037 | -103.493
8.0407 | -103.4942
8.0444 | -103.4943
8.0481 | -103.4831
8.0519 | -103.4734
8.0556 | -103.4596
8.0593 | -103.4449
8.063 | -103.4336
8.0667 | -103.4233
8.0704 | -103.4114
8.0741 | -103.398
8.0778 | -103.3817
8.0815 | -103.3729
8.0852 | -103.3674
8.0889 | -103.362
8.0926 | -103.3558
8.0963 | -103.3481
8.1 | -103.3376
8.1037 | -103.3308
8.1074 | -103.3284
8.1111 | -103.3232
8.1148 | -103.3132
8.1185 | -103.3085
8.1222 | -103.3106
8.1259 | -103.3117
8.1296 | -103.3161
8.1333 | -103.3201
8.137 | -103.3288
8.1407 | -103.3384
8.1444 | -103.3457
8.1481 | -103.3523
8.1519 | -103.3625
8.1556 | -103.3722
8.1593 | -103.3825
8.163 | -103.3911
8.1667 | -103.3944
8.1704 | -103.398
8.1741 | -103.3998
8.1778 | -103.4039
8.1815 | -103.4031
8.1852 | -103.3979
8.1889 | -103.394
8.1926 | -103.3916
8.1963 | -103.3933
8.2 | -103.3969
8.2037 | -103.3956
8.2074 | -103.3943
8.2111 | -103.3976
8.2148 | -103.4016
8.2185 | -103.406
8.2222 | -103.4157
8.2259 | -103.4229
8.2296 | -103.4261
8.2333 | -103.4276
8.237 | -103.4254
8.2407 | -103.4183
8.2444 | -103.4031
8.2481 | -103.3872
8.2519 | -103.3645
8.2556 | -103.3403
8.2593 | -103.3146
8.263 | -103.2934
8.2667 | -103.275
8.2704 | -103.2615
8.2741 | -103.2528
8.2778 | -103.25
8.2815 | -103.2565
8.2852 | -103.2701
8.2889 | -103.2936
8.2926 | -103.3244
8.2963 | -103.3598
8.3 | -103.4002
8.3037 | -103.4484
8.3074 | -103.4989
8.3111 | -103.5489
8.3148 | -103.5983
8.3185 | -103.6477
8.3222 | -103.6954
8.3259 | -103.7396
8.3296 | -103.7833
8.3333 | -103.8198
8.337 | -103.8535
8.3407 | -103.8832
8.3444 | -103.9053
8.3481 | -103.9248
8.3519 | -103.939
8.3556 | -103.9526
8.3593 | -103.9601
8.363 | -103.9614
8.3667 | -103.9619
8.3704 | -103.9582
8.3741 | -103.9577
8.3778 | -103.9546
8.3815 | -103.9485
8.3852 | -103.9386
8.3889 | -103.9353
8.3926 | -103.9399
8.3963 | -103.9488
8.4 | -103.9628
8.4037 | -103.9826
8.4074 | -104.006
8.4111 | -104.0233
8.4148 | -104.0348
8.4185 | -104.046
8.4222 | -104.0555
8.4259 | -104.0555
8.4296 | -104.0574
8.4333 | -104.0528
8.437 | -104.0402
8.4407 | -104.0207
8.4444 | -103.9986
8.4481 | -103.9727
8.4519 | -103.94
8.4556 | -103.9145
8.4593 | -103.8915
8.463 | -103.8685
8.4667 | -103.8439
8.4704 | -103.8258
8.4741 | -103.8127
8.4778 | -103.8024
8.4815 | -103.797
8.4852 | -103.7933
8.4889 | -103.7905
8.4926 | -103.7859
8.4963 | -103.776
8.5 | -103.7658
8.5037 | -103.7517
8.5074 | -103.7359
8.5111 | -103.7163
8.5148 | -103.6929
8.5185 | -103.6714
8.5222 | -103.6477
8.5259 | -103.6194
8.5296 | -103.5882
8.5333 | -103.563
8.537 | -103.5371
8.5407 | -103.5074
8.5444 | -103.4736
8.5481 | -103.4422
8.5519 | -103.4092
8.5556 | -103.3733
8.5593 | -103.3445
8.563 | -103.3186
8.5667 | -103.2963
8.5704 | -103.2796
8.5741 | -103.2684
8.5778 | -103.2597
8.5815 | -103.2516
8.5852 | -103.2466
8.5889 | -103.2474
8.5926 | -103.2541
8.5963 | -103.2599
8.6 | -103.2671
8.6037 | -103.2738
8.6074 | -103.2833
8.6111 | -103.2958
8.6148 | -103.3069
8.6185 | -103.3161
8.6222 | -103.3294
8.6259 | -103.339
8.6296 | -103.3488
8.6333 | -103.3591
8.637 | -103.3691
8.6407 | -103.3792
8.6444 | -103.3886
8.6481 | -103.4029
8.6519 | -103.4196
8.6556 | -103.4409
8.6593 | -103.4618
8.663 | -103.4814
8.6667 | -103.5069
8.6704 | -103.5371
8.6741 | -103.5709
8.6778 | -103.5999
8.6815 | -103.6317
8.6852 | -103.6652
8.6889 | -103.6935
8.6926 | -103.7254
8.6963 | -103.7531
8.7 | -103.7762
8.7037 | -103.7952
8.7074 | -103.8169
8.7111 | -103.8353
8.7148 | -103.8498
8.7185 | -103.86
8.7222 | -103.8628
8.7259 | -103.8554
8.7296 | -103.8489
8.7333 | -103.8449
8.737 | -103.8333
8.7407 | -103.815
8.7444 | -103.8033
8.7481 | -103.7928
8.7519 | -103.7822
8.7556 | -103.7726
8.7593 | -103.759
8.763 | -103.7438
8.7667 | -103.7317
8.7704 | -103.7214
8.7741 | -103.7103
8.7778 | -103.6997
8.7815 | -103.6914
8.7852 | -103.6782
8.7889 | -103.6627
8.7926 | -103.6499
8.7963 | -103.6366
8.8 | -103.6254
8.8037 | -103.6153
8.8074 | -103.6119
8.8111 | -103.6082
8.8148 | -103.6048
8.8185 | -103.6031
8.8222 | -103.5976
8.8259 | -103.6027
8.8296 | -103.6101
8.8333 | -103.6166
8.837 | -103.6202
8.8407 | -103.6206
8.8444 | -103.62
8.8481 | -103.6195
8.8519 | -103.616
8.8556 | -103.6105
8.8593 | -103.5975
8.863 | -103.5803
8.8667 | -103.5652
8.8704 | -103.5498
8.8741 | -103.5292
8.8778 | -103.5079
8.8815 | -103.4864
8.8852 | -103.471
8.8889 | -103.4598
8.8926 | -103.4525
8.8963 | -103.4479
8.9 | -103.4458
8.9037 | -103.4491
8.9074 | -103.4544
8.9111 | -103.4623
8.9148 | -103.4769
8.9185 | -103.4949
8.9222 | -103.513
8.9259 | -103.5342
8.9296 | -103.556
8.9333 | -103.5744
8.937 | -103.5924
8.9407 | -103.6131
8.9444 | -103.6319
8.9481 | -103.6491
8.9519 | -103.6659
8.9556 | -103.675
8.9593 | -103.682
8.963 | -103.6902
8.9667 | -103.6939
8.9704 | -103.6939
8.9741 | -103.6904
8.9778 | -103.6815
8.9815 | -103.6743
8.9852 | -103.665
8.9889 | -103.6501
8.9926 | -103.636
8.9963 | -103.6202
9 | -103.6085
9.0037 | -103.5996
9.0074 | -103.5928
9.0111 | -103.5895
9.0148 | -103.5863
9.0185 | -103.5859
9.0222 | -103.5884
9.0259 | -103.5925
9.0296 | -103.6032
9.0333 | -103.6119
9.037 | -103.6195
9.0407 | -103.6277
9.0444 | -103.6369
9.0481 | -103.6449
9.0519 | -103.652
9.0556 | -103.6557
9.0593 | -103.6576
9.063 | -103.6657
9.0667 | -103.6777
9.0704 | -103.6847
9.0741 | -103.6922
9.0778 | -103.6991
9.0815 | -103.7047
9.0852 | -103.7093
9.0889 | -103.7144
9.0926 | -103.7188
9.0963 | -103.7216
9.1 | -103.7241
9.1037 | -103.7288
9.1074 | -103.7312
9.1111 | -103.7342
9.1148 | -103.7396
9.1185 | -103.7475
9.1222 | -103.7564
9.1259 | -103.7654
9.1296 | -103.7744
9.1333 | -103.7827
9.137 | -103.7906
9.1407 | -103.801
9.1444 | -103.8052
9.1481 | -103.8081
9.1519 | -103.8075
9.1556 | -103.8012
9.1593 | -103.7952
9.163 | -103.7852
9.1667 | -103.7752
9.1704 | -103.7687
9.1741 | -103.7586
9.1778 | -103.7438
9.1815 | -103.7302
9.1852 | -103.7185
9.1889 | -103.7055
9.1926 | -103.6972
9.1963 | -103.6852
9.2 | -103.6708
9.2037 | -103.6658
9.2074 | -103.6616
9.2111 | -103.661
9.2148 | -103.6682
9.2185 | -103.6788
9.2222 | -103.6926
9.2259 | -103.7088
9.2296 | -103.7258
9.2333 | -103.7465
9.237 | -103.769
9.2407 | -103.7937
9.2444 | -103.82
9.2481 | -103.8423
9.2519 | -103.859
9.2556 | -103.8765
9.2593 | -103.8869
9.263 | -103.8909
9.2667 | -103.8936
9.2704 | -103.891
9.2741 | -103.8847
9.2778 | -103.8767
9.2815 | -103.8694
9.2852 | -103.8602
9.2889 | -103.8501
9.2926 | -103.8408
9.2963 | -103.8311
9.3 | -103.8243
9.3037 | -103.82
9.3074 | -103.8164
9.3111 | -103.8173
9.3148 | -103.8195
9.3185 | -103.8238
9.3222 | -103.832
9.3259 | -103.8407
9.3296 | -103.8499
9.3333 | -103.8616
9.337 | -103.8734
9.3407 | -103.8819
9.3444 | -103.897
9.3481 | -103.9095
9.3519 | -103.9216
9.3556 | -103.9371
9.3593 | -103.9476
9.363 | -103.9536
9.3667 | -103.9616
9.3704 | -103.9674
9.3741 | -103.9707
9.3778 | -103.9711
9.3815 | -103.9676
9.3852 | -103.9639
9.3889 | -103.9579
9.3926 | -103.9473
9.3963 | -103.938
9.4 | -103.9201
9.4037 | -103.8994
9.4074 | -103.879
9.4111 | -103.8578
9.4148 | -103.8371
9.4185 | -103.8175
9.4222 | -103.8011
9.4259 | -103.788
9.4296 | -103.7759
9.4333 | -103.7694
9.437 | -103.7695
9.4407 | -103.7746
9.4444 | -103.7855
9.4481 | -103.8055
9.4519 | -103.8337
9.4556 | -103.8635
9.4593 | -103.8979
9.463 | -103.9419
9.4667 | -103.9892
9.4704 | -104.0359
9.4741 | -104.0804
9.4778 | -104.1231
9.4815 | -104.1648
9.4852 | -104.2082
9.4889 | -104.2447
9.4926 | -104.2694
9.4963 | -104.2942
9.5 | -104.3148
9.5037 | -104.3246
9.5074 | -104.3278
9.5111 | -104.325
9.5148 | -104.3117
9.5185 | -104.2948
9.5222 | -104.2766
9.5259 | -104.2574
9.5296 | -104.237
9.5333 | -104.2152
9.537 | -104.1921
9.5407 | -104.1678
9.5444 | -104.1435
9.5481 | -104.1212
9.5519 | -104.1023
9.5556 | -104.0821
9.5593 | -104.0623
9.563 | -104.0454
9.5667 | -104.0258
9.5704 | -104.0011
9.5741 | -103.9784
9.5778 | -103.9565
9.5815 | -103.9344
9.5852 | -103.912
9.5889 | -103.8935
9.5926 | -103.8788
9.5963 | -103.8602
9.6 | -103.8401
9.6037 | -103.8188
9.6074 | -103.8072
9.6111 | -103.8017
9.6148 | -103.8049
9.6185 | -103.8005
9.6222 | -103.7961
9.6259 | -103.7961
9.6296 | -103.8006
9.6333 | -103.8148
9.637 | -103.8354
9.6407 | -103.8567
9.6444 | -103.8821
9.6481 | -103.9082
9.6519 | -103.9372
9.6556 | -103.9697
9.6593 | -104.0049
9.663 | -104.037
9.6667 | -104.0712
9.6704 | -104.1047
9.6741 | -104.1357
9.6778 | -104.1689
9.6815 | -104.196
9.6852 | -104.2168
9.6889 | -104.2335
9.6926 | -104.2489
9.6963 | -104.2662
9.7 | -104.2805
9.7037 | -104.2895
9.7074 | -104.2998
9.7111 | -104.3095
9.7148 | -104.3116
9.7185 | -104.3089
9.7222 | -104.2986
9.7259 | -104.2876
9.7296 | -104.2799
9.7333 | -104.2705
9.737 | -104.2649
9.7407 | -104.2602
9.7444 | -104.2535
9.7481 | -104.2377
9.7519 | -104.2209
9.7556 | -104.206
9.7593 | -104.1957
9.763 | -104.1845
9.7667 | -104.1694
9.7704 | -104.1546
9.7741 | -104.1406
9.7778 | -104.1208
9.7815 | -104.0991
9.7852 | -104.0804
9.7889 | -104.063
9.7926 | -104.0467
9.7963 | -104.0334
9.8 | -104.0217
9.8037 | -104.0089
9.8074 | -104.0004
9.8111 | -103.9933
9.8148 | -103.9835
9.8185 | -103.975
9.8222 | -103.9689
9.8259 | -103.9646
9.8296 | -103.9649
9.8333 | -103.9646
9.837 | -103.9613
9.8407 | -103.9613
9.8444 | -103.9677
9.8481 | -103.9742
9.8519 | -103.9825
9.8556 | -103.9937
9.8593 | -104.0113
9.863 | -104.0316
9.8667 | -104.0571
9.8704 | -104.079
9.8741 | -104.0938
9.8778 | -104.106
9.8815 | -104.1203
9.8852 | -104.1399
9.8889 | -104.162
9.8926 | -104.1816
9.8963 | -104.1969
9.9 | -104.2126
9.9037 | -104.2224
9.9074 | -104.2304
9.9111 | -104.2386
9.9148 | -104.2438
9.9185 | -104.2475
9.9222 | -104.2548
9.9259 | -104.2566
9.9296 | -104.2509
9.9333 | -104.2467
9.937 | -104.2448
9.9407 | -104.2418
9.9444 | -104.2377
9.9481 | -104.2329
9.9519 | -104.2253
9.9556 | -104.2152
9.9593 | -104.2118
9.963 | -104.2112
9.9667 | -104.2094
9.9704 | -104.2088
9.9741 | -104.2105
9.9778 | -104.2157
9.9815 | -104.2227
9.9852 | -104.2322
9.9889 | -104.2446
9.9926 | -104.2604
9.9963 | -104.2746
10 | -104.2921
10.0037 | -104.3111
10.0074 | -104.3325
10.0111 | -104.3519
10.0148 | -104.3731
10.0185 | -104.3981
10.0222 | -104.4268
10.0259 | -104.4508
10.0296 | -104.4676
10.0333 | -104.4869
10.037 | -104.5079
10.0407 | -104.5261
10.0444 | -104.5377
10.0481 | -104.5442
10.0519 | -104.5515
10.0556 | -104.5562
10.0593 | -104.5593
10.063 | -104.5615
10.0667 | -104.5593
10.0704 | -104.5548
10.0741 | -104.551
10.0778 | -104.5394
10.0815 | -104.5205
10.0852 | -104.4998
10.0889 | -104.4821
10.0926 | -104.4627
10.0963 | -104.4409
10.1 | -104.4228
10.1037 | -104.4031
10.1074 | -104.3802
10.1111 | -104.3595
10.1148 | -104.3428
10.1185 | -104.3229
10.1222 | -104.3048
10.1259 | -104.2918
10.1296 | -104.2799
10.1333 | -104.2693
10.137 | -104.263
10.1407 | -104.2608
10.1444 | -104.2551
10.1481 | -104.254
10.1519 | -104.2648
10.1556 | -104.2604
10.1593 | -104.2638
10.163 | -104.2676
10.1667 | -104.2712
10.1704 | -104.2768
10.1741 | -104.2812
10.1778 | -104.2832
10.1815 | -104.2861
10.1852 | -104.2897
10.1889 | -104.2895
10.1926 | -104.2878
10.1963 | -104.2861
10.2 | -104.2772
10.2037 | -104.2642
10.2074 | -104.252
10.2111 | -104.2363
10.2148 | -104.2207
10.2185 | -104.2032
10.2222 | -104.181
10.2259 | -104.1624
10.2296 | -104.15
10.2333 | -104.1375
10.237 | -104.1203
10.2407 | -104.1088
10.2444 | -104.1013
10.2481 | -104.092
10.2519 | -104.0858
10.2556 | -104.0863
10.2593 | -104.0875
10.263 | -104.0914
10.2667 | -104.0999
10.2704 | -104.1061
10.2741 | -104.1125
10.2778 | -104.1194
10.2815 | -104.1259
10.2852 | -104.1371
10.2889 | -104.151
10.2926 | -104.1672
10.2963 | -104.1811
10.3 | -104.1977
10.3037 | -104.2153
10.3074 | -104.237
10.3111 | -104.2507
10.3148 | -104.2654
10.3185 | -104.2821
10.3222 | -104.2947
10.3259 | -104.3034
10.3296 | -104.3108
10.3333 | -104.3173
10.337 | -104.319
10.3407 | -104.3245
10.3444 | -104.3312
10.3481 | -104.336
10.3519 | -104.3395
10.3556 | -104.3414
10.3593 | -104.3413
10.363 | -104.3464
10.3667 | -104.3541
10.3704 | -104.3631
10.3741 | -104.3756
10.3778 | -104.3861
10.3815 | -104.3952
10.3852 | -104.4059
10.3889 | -104.4208
10.3926 | -104.4354
10.3963 | -104.4477
10.4 | -104.4597
10.4037 | -104.4749
10.4074 | -104.488
10.4111 | -104.5008
10.4148 | -104.5121
10.4185 | -104.5206
10.4222 | -104.5315
10.4259 | -104.5394
10.4296 | -104.5438
10.4333 | -104.546
10.437 | -104.5412
10.4407 | -104.5336
10.4444 | -104.5257
10.4481 | -104.5136
10.4519 | -104.4921
10.4556 | -104.4654
10.4593 | -104.4322
10.463 | -104.3945
10.4667 | -104.3512
10.4704 | -104.2982
10.4741 | -104.2423
10.4778 | -104.1885
10.4815 | -104.1366
10.4852 | -104.0817
10.4889 | -104.0327
10.4926 | -103.9804
10.4963 | -103.9285
10.5 | -103.8849
10.5037 | -103.8466
10.5074 | -103.817
10.5111 | -103.7934
10.5148 | -103.7766
10.5185 | -103.7661
10.5222 | -103.7634
10.5259 | -103.769
10.5296 | -103.7791
10.5333 | -103.7935
10.537 | -103.8119
10.5407 | -103.8362
10.5444 | -103.8656
10.5481 | -103.9
10.5519 | -103.9407
10.5556 | -103.977
10.5593 | -104.0103
10.563 | -104.0423
10.5667 | -104.0759
10.5704 | -104.1116
10.5741 | -104.142
10.5778 | -104.1676
10.5815 | -104.1934
10.5852 | -104.2151
10.5889 | -104.2301
10.5926 | -104.2405
10.5963 | -104.2475
10.6 | -104.2537
10.6037 | -104.263
10.6074 | -104.2712
10.6111 | -104.2773
10.6148 | -104.2847
10.6185 | -104.2916
10.6222 | -104.3012
10.6259 | -104.3144
10.6296 | -104.3289
10.6333 | -104.3437
10.637 | -104.3622
10.6407 | -104.3845
10.6444 | -104.4159
10.6481 | -104.4522
10.6519 | -104.4863
10.6556 | -104.5295
10.6593 | -104.5709
10.663 | -104.6117
10.6667 | -104.6543
10.6704 | -104.7002
10.6741 | -104.7453
10.6778 | -104.7864
10.6815 | -104.8288
10.6852 | -104.8598
10.6889 | -104.8952
10.6926 | -104.9289
10.6963 | -104.9566
10.7 | -104.9784
10.7037 | -104.996
10.7074 | -105.0045
10.7111 | -105.0149
10.7148 | -105.0253
10.7185 | -105.033
10.7222 | -105.0282
10.7259 | -105.0169
10.7296 | -105.0063
10.7333 | -104.9906
10.737 | -104.9684
10.7407 | -104.9438
10.7444 | -104.9166
10.7481 | -104.8841
10.7519 | -104.8462
10.7556 | -104.8074
10.7593 | -104.7669
10.763 | -104.7227
10.7667 | -104.6771
10.7704 | -104.6299
10.7741 | -104.5872
10.7778 | -104.5448
10.7815 | -104.4976
10.7852 | -104.4536
10.7889 | -104.4166
10.7926 | -104.3849
10.7963 | -104.3552
10.8 | -104.3295
10.8037 | -104.3094
10.8074 | -104.2998
10.8111 | -104.2909
10.8148 | -104.2875
10.8185 | -104.2921
10.8222 | -104.3017
10.8259 | -104.3132
10.8296 | -104.3338
10.8333 | -104.3592
10.837 | -104.3861
10.8407 | -104.4144
10.8444 | -104.4471
10.8481 | -104.481
10.8519 | -104.5196
10.8556 | -104.5601
10.8593 | -104.5955
10.863 | -104.6273
10.8667 | -104.6594
10.8704 | -104.689
10.8741 | -104.7167
10.8778 | -104.7415
10.8815 | -104.7641
10.8852 | -104.7831
10.8889 | -104.796
10.8926 | -104.8083
10.8963 | -104.8201
10.9 | -104.8258
10.9037 | -104.8289
10.9074 | -104.8307
10.9111 | -104.8291
10.9148 | -104.8233
10.9185 | -104.8143
10.9222 | -104.804
10.9259 | -104.7902
10.9296 | -104.7748
10.9333 | -104.7526
10.937 | -104.7315
10.9407 | -104.7145
10.9444 | -104.6941
10.9481 | -104.6725
10.9519 | -104.6501
10.9556 | -104.6294
10.9593 | -104.614
10.963 | -104.6044
10.9667 | -104.5912
10.9704 | -104.579
10.9741 | -104.5724
10.9778 | -104.5662
10.9815 | -104.5611
10.9852 | -104.5615
10.9889 | -104.557
10.9926 | -104.5521
10.9963 | -104.5503
11 | -104.5498
11.0037 | -104.5531
11.0074 | -104.5576
11.0111 | -104.5592
11.0148 | -104.5594
11.0185 | -104.5634
11.0222 | -104.5705
11.0259 | -104.576
11.0296 | -104.5837
11.0333 | -104.5902
11.037 | -104.5952
11.0407 | -104.5982
11.0444 | -104.5991
11.0481 | -104.6034
11.0519 | -104.6078
11.0556 | -104.6104
11.0593 | -104.6169
11.063 | -104.623
11.0667 | -104.6273
11.0704 | -104.6371
11.0741 | -104.6472
11.0778 | -104.6571
11.0815 | -104.6657
11.0852 | -104.6758
11.0889 | -104.6844
11.0926 | -104.6911
11.0963 | -104.7039
11.1 | -104.7147
11.1037 | -104.7263
11.1074 | -104.729
11.1111 | -104.7301
11.1148 | -104.7316
11.1185 | -104.7312
11.1222 | -104.725
11.1259 | -104.7217
11.1296 | -104.7254
11.1333 | -104.7261
11.137 | -104.722
11.1407 | -104.7161
11.1444 | -104.7085
11.1481 | -104.7005
11.1519 | -104.6911
11.1556 | -104.6829
11.1593 | -104.6732
11.163 | -104.657
11.1667 | -104.6405
11.1704 | -104.623
11.1741 | -104.6052
11.1778 | -104.5876
11.1815 | -104.5687
11.1852 | -104.5476
11.1889 | -104.5241
11.1926 | -104.505
11.1963 | -104.4905
11.2 | -104.4751
11.2037 | -104.4625
11.2074 | -104.448
11.2111 | -104.4318
11.2148 | -104.4162
11.2185 | -104.4047
11.2222 | -104.3918
11.2259 | -104.3823
11.2296 | -104.3731
11.2333 | -104.362
11.237 | -104.3534
11.2407 | -104.3475
11.2444 | -104.3449
11.2481 | -104.3398
11.2519 | -104.3329
11.2556 | -104.3288
11.2593 | -104.3216
11.263 | -104.3146
11.2667 | -104.3082
11.2704 | -104.3051
11.2741 | -104.3003
11.2778 | -104.2981
11.2815 | -104.2971
11.2852 | -104.2974
11.2889 | -104.2979
11.2926 | -104.2942
11.2963 | -104.2909
11.3 | -104.2943
11.3037 | -104.3013
11.3074 | -104.3093
11.3111 | -104.3175
11.3148 | -104.329
11.3185 | -104.3378
11.3222 | -104.3494
11.3259 | -104.3654
11.3296 | -104.385
11.3333 | -104.4046
11.337 | -104.4266
11.3407 | -104.453
11.3444 | -104.4812
11.3481 | -104.5114
11.3519 | -104.5409
11.3556 | -104.5687
11.3593 | -104.5937
11.363 | -104.6265
11.3667 | -104.6575
11.3704 | -104.6828
11.3741 | -104.7081
11.3778 | -104.7351
11.3815 | -104.759
11.3852 | -104.7794
11.3889 | -104.7924
11.3926 | -104.8028
11.3963 | -104.8103
11.4 | -104.8132
11.4037 | -104.8118
11.4074 | -104.809
11.4111 | -104.8055
11.4148 | -104.8
11.4185 | -104.7925
11.4222 | -104.782
11.4259 | -104.7705
11.4296 | -104.7586
11.4333 | -104.745
11.437 | -104.7303
11.4407 | -104.7168
11.4444 | -104.6927
11.4481 | -104.6636
11.4519 | -104.6379
11.4556 | -104.6123
11.4593 | -104.5862
11.463 | -104.5616
11.4667 | -104.5361
11.4704 | -104.5098
11.4741 | -104.4846
11.4778 | -104.4596
11.4815 | -104.4299
11.4852 | -104.4029
11.4889 | -104.3751
11.4926 | -104.3485
11.4963 | -104.3219
11.5 | -104.2998
11.5037 | -104.2768
11.5074 | -104.2502
11.5111 | -104.2254
11.5148 | -104.2068
11.5185 | -104.1967
11.5222 | -104.1888
11.5259 | -104.1841
11.5296 | -104.1819
11.5333 | -104.1841
11.537 | -104.194
11.5407 | -104.2104
11.5444 | -104.2303
11.5481 | -104.254
11.5519 | -104.283
11.5556 | -104.3167
11.5593 | -104.3504
11.563 | -104.3863
11.5667 | -104.4195
11.5704 | -104.4533
11.5741 | -104.4857
11.5778 | -104.5146
11.5815 | -104.5402
11.5852 | -104.5624
11.5889 | -104.5818
11.5926 | -104.5965
11.5963 | -104.6085
11.6 | -104.6175
11.6037 | -104.6253
11.6074 | -104.6297
11.6111 | -104.6369
11.6148 | -104.6416
11.6185 | -104.6434
11.6222 | -104.6448
11.6259 | -104.6461
11.6296 | -104.6458
11.6333 | -104.6443
11.637 | -104.6458
11.6407 | -104.6493
11.6444 | -104.6602
11.6481 | -104.6707
11.6519 | -104.6786
11.6556 | -104.689
11.6593 | -104.703
11.663 | -104.7169
11.6667 | -104.7329
11.6704 | -104.7553
11.6741 | -104.7813
11.6778 | -104.8079
11.6815 | -104.8304
11.6852 | -104.8485
11.6889 | -104.8688
11.6926 | -104.8899
11.6963 | -104.9066
11.7 | -104.9235
11.7037 | -104.9421
11.7074 | -104.9544
11.7111 | -104.9629
11.7148 | -104.972
11.7185 | -104.9795
11.7222 | -104.9825
11.7259 | -104.9823
11.7296 | -104.9803
11.7333 | -104.9841
11.737 | -104.9858
11.7407 | -104.9835
11.7444 | -104.9814
11.7481 | -104.9811
11.7519 | -104.9807
11.7556 | -104.9807
11.7593 | -104.9781
11.763 | -104.9747
11.7667 | -104.9741
11.7704 | -104.9736
11.7741 | -104.9718
11.7778 | -104.9712
11.7815 | -104.9686
11.7852 | -104.9647
11.7889 | -104.96
11.7926 | -104.9522
11.7963 | -104.9434
11.8 | -104.9362
11.8037 | -104.9303
11.8074 | -104.9189
11.8111 | -104.9051
11.8148 | -104.8906
11.8185 | -104.8816
11.8222 | -104.8714
11.8259 | -104.8595
11.8296 | -104.8417
11.8333 | -104.8242
11.837 | -104.8044
11.8407 | -104.7812
11.8444 | -104.7548
11.8481 | -104.7288
11.8519 | -104.7043
11.8556 | -104.6831
11.8593 | -104.6598
11.863 | -104.6334
11.8667 | -104.6085
11.8704 | -104.5777
11.8741 | -104.5374
11.8778 | -104.497
11.8815 | -104.4577
11.8852 | -104.421
11.8889 | -104.3821
11.8926 | -104.345
11.8963 | -104.3094
11.9 | -104.2754
11.9037 | -104.2394
11.9074 | -104.2054
11.9111 | -104.1748
11.9148 | -104.1471
11.9185 | -104.1231
11.9222 | -104.1026
11.9259 | -104.0867
11.9296 | -104.0746
11.9333 | -104.0664
11.937 | -104.0616
11.9407 | -104.0603
11.9444 | -104.0646
11.9481 | -104.0778
11.9519 | -104.0942
11.9556 | -104.1159
11.9593 | -104.1417
11.963 | -104.1694
11.9667 | -104.206
11.9704 | -104.2486
11.9741 | -104.2927
11.9778 | -104.3376
11.9815 | -104.3864
11.9852 | -104.4362
11.9889 | -104.4858
11.9926 | -104.5327
11.9963 | -104.5799
12 | -104.6252
12.0037 | -104.6696
12.0074 | -104.711
12.0111 | -104.7473
12.0148 | -104.7801
12.0185 | -104.8097
12.0222 | -104.8303
12.0259 | -104.845
12.0296 | -104.8524
12.0333 | -104.8545
12.037 | -104.8492
12.0407 | -104.84
12.0444 | -104.8317
12.0481 | -104.8257
12.0519 | -104.8127
12.0556 | -104.7921
12.0593 | -104.7661
12.063 | -104.7443
12.0667 | -104.7199
12.0704 | -104.6921
12.0741 | -104.6646
12.0778 | -104.6414
12.0815 | -104.6186
12.0852 | -104.5983
12.0889 | -104.5808
12.0926 | -104.5629
12.0963 | -104.5469
12.1 | -104.5335
12.1037 | -104.5274
12.1074 | -104.5226
12.1111 | -104.5196
12.1148 | -104.5205
12.1185 | -104.5247
12.1222 | -104.5296
12.1259 | -104.5379
12.1296 | -104.545
12.1333 | -104.5508
12.137 | -104.5542
12.1407 | -104.5539
12.1444 | -104.5536
12.1481 | -104.5562
12.1519 | -104.5584
12.1556 | -104.5624
12.1593 | -104.5682
12.163 | -104.5702
12.1667 | -104.57
12.1704 | -104.5689
12.1741 | -104.5664
12.1778 | -104.5631
12.1815 | -104.5603
12.1852 | -104.5598
12.1889 | -104.5602
12.1926 | -104.5601
12.1963 | -104.5607
12.2 | -104.5632
12.2037 | -104.5678
12.2074 | -104.5761
12.2111 | -104.5893
12.2148 | -104.6034
12.2185 | -104.6191
12.2222 | -104.6359
12.2259 | -104.6579
12.2296 | -104.6848
12.2333 | -104.7118
12.237 | -104.738
12.2407 | -104.7668
12.2444 | -104.798
12.2481 | -104.8269
12.2519 | -104.8558
12.2556 | -104.8866
12.2593 | -104.9182
12.263 | -104.9498
12.2667 | -104.9811
12.2704 | -105.0092
12.2741 | -105.0317
12.2778 | -105.0462
12.2815 | -105.061
12.2852 | -105.0736
12.2889 | -105.0643
12.2926 | -105.0687
12.2963 | -105.0681
12.3 | -105.0621
12.3037 | -105.0576
12.3074 | -105.0543
12.3111 | -105.0499
12.3148 | -105.0439
12.3185 | -105.0351
12.3222 | -105.0208
12.3259 | -105.0039
12.3296 | -104.9876
12.3333 | -104.9702
12.337 | -104.9546
12.3407 | -104.9403
12.3444 | -104.9282
12.3481 | -104.9158
12.3519 | -104.9054
12.3556 | -104.8962
12.3593 | -104.8875
12.363 | -104.8807
12.3667 | -104.8741
12.3704 | -104.8678
12.3741 | -104.8564
12.3778 | -104.8446
12.3815 | -104.8365
12.3852 | -104.8313
12.3889 | -104.8267
12.3926 | -104.8229
12.3963 | -104.8165
12.4 | -104.8118
12.4037 | -104.8054
12.4074 | -104.7985
12.4111 | -104.7939
12.4148 | -104.7857
12.4185 | -104.7771
12.4222 | -104.7716
12.4259 | -104.7646
12.4296 | -104.7522
12.4333 | -104.7362
12.437 | -104.7189
12.4407 | -104.7015
12.4444 | -104.6833
12.4481 | -104.6661
12.4519 | -104.6473
12.4556 | -104.6282
12.4593 | -104.6067
12.463 | -104.5826
12.4667 | -104.5592
12.4704 | -104.527
12.4741 | -104.4918
12.4778 | -104.4613
12.4815 | -104.4308
12.4852 | -104.3977
12.4889 | -104.3681
12.4926 | -104.3361
12.4963 | -104.2991
12.5 | -104.2619
12.5037 | -104.2281
12.5074 | -104.1968
12.5111 | -104.1663
12.5148 | -104.1361
12.5185 | -104.1064
12.5222 | -104.0795
12.5259 | -104.0562
12.5296 | -104.0372
12.5333 | -104.0191
12.537 | -104.0041
12.5407 | -103.9943
12.5444 | -103.9901
12.5481 | -103.9885
12.5519 | -103.9904
12.5556 | -104.0002
12.5593 | -104.0128
12.563 | -104.0238
12.5667 | -104.0376
12.5704 | -104.0553
12.5741 | -104.074
12.5778 | -104.0996
12.5815 | -104.124
12.5852 | -104.1489
12.5889 | -104.1726
12.5926 | -104.1934
12.5963 | -104.2147
12.6 | -104.231
12.6037 | -104.242
12.6074 | -104.2543
12.6111 | -104.2681
12.6148 | -104.2824
12.6185 | -104.2909
12.6222 | -104.2973
12.6259 | -104.304
12.6296 | -104.3081
12.6333 | -104.3123
12.637 | -104.3172
12.6407 | -104.3183
12.6444 | -104.3213
12.6481 | -104.327
12.6519 | -104.3369
12.6556 | -104.346
12.6593 | -104.3537
12.663 | -104.3611
12.6667 | -104.3676
12.6704 | -104.371
12.6741 | -104.3757
12.6778 | -104.3854
12.6815 | -104.3971
12.6852 | -104.4073
12.6889 | -104.4132
12.6926 | -104.4193
12.6963 | -104.4244
12.7 | -104.4283
12.7037 | -104.4303
12.7074 | -104.428
12.7111 | -104.4235
12.7148 | -104.4205
12.7185 | -104.4167
12.7222 | -104.407
12.7259 | -104.3906
12.7296 | -104.3738
12.7333 | -104.361
12.737 | -104.3502
12.7407 | -104.3369
12.7444 | -104.3234
12.7481 | -104.3067
12.7519 | -104.2871
12.7556 | -104.2683
12.7593 | -104.2514
12.763 | -104.2406
12.7667 | -104.2282
12.7704 | -104.2161
12.7741 | -104.207
12.7778 | -104.2
12.7815 | -104.1982
12.7852 | -104.1952
12.7889 | -104.1912
12.7926 | -104.1891
12.7963 | -104.1905
12.8 | -104.1969
12.8037 | -104.2047
12.8074 | -104.2144
12.8111 | -104.2298
12.8148 | -104.2491
12.8185 | -104.2711
12.8222 | -104.3124
12.8259 | -104.343
12.8296 | -104.3752
12.8333 | -104.4071
12.837 | -104.4458
12.8407 | -104.4881
12.8444 | -104.5267
12.8481 | -104.5669
12.8519 | -104.6119
12.8556 | -104.6583
12.8593 | -104.7032
12.863 | -104.7428
12.8667 | -104.7787
12.8704 | -104.8115
12.8741 | -104.8443
12.8778 | -104.8713
12.8815 | -104.893
12.8852 | -104.9064
12.8889 | -104.9185
12.8926 | -104.9283
12.8963 | -104.9309
12.9 | -104.926
12.9037 | -104.9153
12.9074 | -104.906
12.9111 | -104.8958
12.9148 | -104.8812
12.9185 | -104.8566
12.9222 | -104.8291
12.9259 | -104.8001
12.9296 | -104.773
12.9333 | -104.7492
12.937 | -104.7263
12.9407 | -104.7039
12.9444 | -104.6809
12.9481 | -104.6618
12.9519 | -104.6474
12.9556 | -104.6361
12.9593 | -104.6295
12.963 | -104.6243
12.9667 | -104.6203
12.9704 | -104.6229
12.9741 | -104.6291
12.9778 | -104.6373
12.9815 | -104.6492
12.9852 | -104.6673
12.9889 | -104.6888
12.9926 | -104.7124
12.9963 | -104.7367
13 | -104.7629
13.0037 | -104.7918
13.0074 | -104.8161
13.0111 | -104.8353
13.0148 | -104.8576
13.0185 | -104.8773
13.0222 | -104.8969
13.0259 | -104.9125
13.0296 | -104.9239
13.0333 | -104.9372
13.037 | -104.9507
13.0407 | -104.9579
13.0444 | -104.9608
13.0481 | -104.9648
13.0519 | -104.968
13.0556 | -104.9647
13.0593 | -104.9582
13.063 | -104.9534
13.0667 | -104.9547
13.0704 | -104.9584
13.0741 | -104.9632
13.0778 | -104.9619
13.0815 | -104.9577
13.0852 | -104.958
13.0889 | -104.9579
13.0926 | -104.9602
13.0963 | -104.961
13.1 | -104.9624
13.1037 | -104.9702
13.1074 | -104.9815
13.1111 | -104.9878
13.1148 | -104.9935
13.1185 | -105.0001
13.1222 | -105.0079
13.1259 | -105.0133
13.1296 | -105.0183
13.1333 | -105.0244
13.137 | -105.0266
13.1407 | -105.0298
13.1444 | -105.0328
13.1481 | -105.0304
13.1519 | -105.028
13.1556 | -105.0255
13.1593 | -105.0203
13.163 | -105.0118
13.1667 | -105.0023
13.1704 | -104.9892
13.1741 | -104.9721
13.1778 | -104.9551
13.1815 | -104.9371
13.1852 | -104.9184
13.1889 | -104.9014
13.1926 | -104.8818
13.1963 | -104.8699
13.2 | -104.8638
13.2037 | -104.8545
13.2074 | -104.8462
13.2111 | -104.8397
13.2148 | -104.8345
13.2185 | -104.8331
13.2222 | -104.8289
13.2259 | -104.8272
13.2296 | -104.8335
13.2333 | -104.8438
13.237 | -104.8568
13.2407 | -104.87
13.2444 | -104.8863
13.2481 | -104.9081
13.2519 | -104.9275
13.2556 | -104.9416
13.2593 | -104.9556
13.263 | -104.9695
13.2667 | -104.9856
13.2704 | -105.0029
13.2741 | -105.0147
13.2778 | -105.0262
13.2815 | -105.0356
13.2852 | -105.043
13.2889 | -105.0512
13.2926 | -105.054
13.2963 | -105.0529
13.3 | -105.0564
13.3037 | -105.0608
13.3074 | -105.0585
13.3111 | -105.0549
13.3148 | -105.0544
13.3185 | -105.0522
13.3222 | -105.0462
13.3259 | -105.0416
13.3296 | -105.0398
13.3333 | -105.0381
13.337 | -105.033
13.3407 | -105.0275
13.3444 | -105.0225
13.3481 | -105.0164
13.3519 | -105.012
13.3556 | -105.0034
13.3593 | -104.995
13.363 | -104.9874
13.3667 | -104.9802
13.3704 | -104.971
13.3741 | -104.9622
13.3778 | -104.9542
13.3815 | -104.9451
13.3852 | -104.934
13.3889 | -104.9232
13.3926 | -104.914
13.3963 | -104.9075
13.4 | -104.9029
13.4037 | -104.8919
13.4074 | -104.8833
13.4111 | -104.8782
13.4148 | -104.8802
13.4185 | -104.8822
13.4222 | -104.8759
13.4259 | -104.8716
13.4296 | -104.8661
13.4333 | -104.8607
13.437 | -104.8588
13.4407 | -104.857
13.4444 | -104.854
13.4481 | -104.8527
13.4519 | -104.8503
13.4556 | -104.8469
13.4593 | -104.8421
13.463 | -104.8364
13.4667 | -104.8294
13.4704 | -104.8202
13.4741 | -104.8098
13.4778 | -104.7956
13.4815 | -104.7779
13.4852 | -104.759
13.4889 | -104.7415
13.4926 | -104.7177
13.4963 | -104.6935
13.5 | -104.6724
13.5037 | -104.6515
13.5074 | -104.6312
13.5111 | -104.6123
13.5148 | -104.5951
13.5185 | -104.5774
13.5222 | -104.5598
13.5259 | -104.5442
13.5296 | -104.5329
13.5333 | -104.5209
13.537 | -104.5082
13.5407 | -104.4983
13.5444 | -104.4915
13.5481 | -104.4886
13.5519 | -104.4871
13.5556 | -104.486
13.5593 | -104.4879
13.563 | -104.4927
13.5667 | -104.5008
13.5704 | -104.5099
13.5741 | -104.5211
13.5778 | -104.5367
13.5815 | -104.5558
13.5852 | -104.5721
13.5889 | -104.5865
13.5926 | -104.6019
13.5963 | -104.6185
13.6 | -104.6401
13.6037 | -104.6622
13.6074 | -104.6823
13.6111 | -104.7
13.6148 | -104.7177
13.6185 | -104.736
13.6222 | -104.7478
13.6259 | -104.7563
13.6296 | -104.7676
13.6333 | -104.7778
13.637 | -104.7856
13.6407 | -104.7925
13.6444 | -104.7969
13.6481 | -104.798
13.6519 | -104.7981
13.6556 | -104.7946
13.6593 | -104.7887
13.663 | -104.7817
13.6667 | -104.7735
13.6704 | -104.761
13.6741 | -104.7483
13.6778 | -104.735
13.6815 | -104.7239
13.6852 | -104.712
13.6889 | -104.7002
13.6926 | -104.6867
13.6963 | -104.6718
13.7 | -104.6562
13.7037 | -104.6411
13.7074 | -104.6264
13.7111 | -104.6131
13.7148 | -104.6075
13.7185 | -104.602
13.7222 | -104.6008
13.7259 | -104.6042
13.7296 | -104.6087
13.7333 | -104.613
13.737 | -104.6176
13.7407 | -104.6258
13.7444 | -104.6393
13.7481 | -104.654
13.7519 | -104.6719
13.7556 | -104.6872
13.7593 | -104.7015
13.763 | -104.7158
13.7667 | -104.732
13.7704 | -104.746
13.7741 | -104.7555
13.7778 | -104.7621
13.7815 | -104.7673
13.7852 | -104.7709
13.7889 | -104.7688
13.7926 | -104.7646
13.7963 | -104.7597
13.8 | -104.7512
13.8037 | -104.7391
13.8074 | -104.7259
13.8111 | -104.7089
13.8148 | -104.6888
13.8185 | -104.6688
13.8222 | -104.6485
13.8259 | -104.6224
13.8296 | -104.5879
13.8333 | -104.5485
13.837 | -104.5114
13.8407 | -104.4784
13.8444 | -104.4472
13.8481 | -104.415
13.8519 | -104.3864
13.8556 | -104.3587
13.8593 | -104.3332
13.863 | -104.31
13.8667 | -104.2927
13.8704 | -104.2771
13.8741 | -104.2644
13.8778 | -104.2516
13.8815 | -104.2427
13.8852 | -104.2376
13.8889 | -104.2562
13.8926 | -104.2601
13.8963 | -104.266
13.9 | -104.2747
13.9037 | -104.2883
13.9074 | -104.3041
13.9111 | -104.3214
13.9148 | -104.3372
13.9185 | -104.3557
13.9222 | -104.3757
13.9259 | -104.3949
13.9296 | -104.4154
13.9333 | -104.4366
13.937 | -104.4544
13.9407 | -104.4671
13.9444 | -104.4752
13.9481 | -104.4837
13.9519 | -104.4921
13.9556 | -104.4994
13.9593 | -104.4927
13.963 | -104.4741
13.9667 | -104.4594
13.9704 | -104.4443
13.9741 | -104.4309
13.9778 | -104.4196
13.9815 | -104.4064
13.9852 | -104.3905
13.9889 | -104.372
13.9926 | -104.349
13.9963 | -104.3223
14 | -104.2932
14.0037 | -104.263
14.0074 | -104.2326
14.0111 | -104.2056
14.0148 | -104.1806
14.0185 | -104.1533
14.0222 | -104.1251
14.0259 | -104.0975
14.0296 | -104.0729
14.0333 | -104.0554
14.037 | -104.0401
14.0407 | -104.0259
14.0444 | -104.0153
14.0481 | -104.0086
14.0519 | -104.0041
14.0556 | -104.0029
14.0593 | -104.0053
14.063 | -104.0122
14.0667 | -104.0224
14.0704 | -104.0383
14.0741 | -104.0592
14.0778 | -104.0855
14.0815 | -104.1136
14.0852 | -104.1459
14.0889 | -104.1787
14.0926 | -104.2144
14.0963 | -104.2521
14.1 | -104.2889
14.1037 | -104.3261
14.1074 | -104.3597
14.1111 | -104.3922
14.1148 | -104.4272
14.1185 | -104.4629
14.1222 | -104.4968
14.1259 | -104.5304
14.1296 | -104.5645
14.1333 | -104.5976
14.137 | -104.6285
14.1407 | -104.6598
14.1444 | -104.6895
14.1481 | -104.7195
14.1519 | -104.7474
14.1556 | -104.772
14.1593 | -104.796
14.163 | -104.8295
14.1667 | -104.8681
14.1704 | -104.8901
14.1741 | -104.9037
14.1778 | -104.9145
14.1815 | -104.9249
14.1852 | -104.9338
14.1889 | -104.9421
14.1926 | -104.9518
14.1963 | -104.9607
14.2 | -104.967
14.2037 | -104.9728
14.2074 | -104.9777
14.2111 | -104.9791
14.2148 | -104.9811
14.2185 | -104.9811
14.2222 | -104.984
14.2259 | -104.9846
14.2296 | -104.9801
14.2333 | -104.9785
14.237 | -104.9736
14.2407 | -104.9679
14.2444 | -104.9621
14.2481 | -104.9573
14.2519 | -104.9555
14.2556 | -104.9557
14.2593 | -104.957
14.263 | -104.9568
14.2667 | -104.9552
14.2704 | -104.9486
14.2741 | -104.9451
14.2778 | -104.9451
14.2815 | -104.9452
14.2852 | -104.9496
14.2889 | -104.9554
14.2926 | -104.9593
14.2963 | -104.9599
14.3 | -104.9616
14.3037 | -104.971
14.3074 | -104.9812
14.3111 | -104.9924
14.3148 | -105.0032
14.3185 | -105.0106
14.3222 | -105.0153
14.3259 | -105.0232
14.3296 | -105.0268
14.3333 | -105.0298
14.337 | -105.0358
14.3407 | -105.0429
14.3444 | -105.0476
14.3481 | -105.0477
14.3519 | -105.0492
14.3556 | -105.051
14.3593 | -105.0554
14.363 | -105.0577
14.3667 | -105.06
14.3704 | -105.0626
14.3741 | -105.0685
14.3778 | -105.0738
14.3815 | -105.0728
14.3852 | -105.0716
14.3889 | -105.0749
14.3926 | -105.0728
14.3963 | -105.0648
14.4 | -105.0576
14.4037 | -105.0514
14.4074 | -105.0424
14.4111 | -105.0323
14.4148 | -105.0203
14.4185 | -105.0017
14.4222 | -104.9831
14.4259 | -104.9466
14.4296 | -104.9226
14.4333 | -104.8981
14.437 | -104.8742
14.4407 | -104.8486
14.4444 | -104.8213
14.4481 | -104.7928
14.4519 | -104.7649
14.4556 | -104.7353
14.4593 | -104.705
14.463 | -104.6782
14.4667 | -104.6544
14.4704 | -104.6351
14.4741 | -104.6128
14.4778 | -104.5956
14.4815 | -104.5772
14.4852 | -104.5588
14.4889 | -104.5416
14.4926 | -104.5243
14.4963 | -104.5098
14.5 | -104.5039
14.5037 | -104.503
14.5074 | -104.503
14.5111 | -104.503
14.5148 | -104.5054
14.5185 | -104.5107
14.5222 | -104.5158
14.5259 | -104.5221
14.5296 | -104.5301
14.5333 | -104.5396
14.537 | -104.5518
14.5407 | -104.5642
14.5444 | -104.5748
14.5481 | -104.5888
14.5519 | -104.6056
14.5556 | -104.6235
14.5593 | -104.6353
14.563 | -104.6463
14.5667 | -104.6552
14.5704 | -104.6627
14.5741 | -104.668
14.5778 | -104.669
14.5815 | -104.6673
14.5852 | -104.6684
14.5889 | -104.6676
14.5926 | -104.6591
14.5963 | -104.6464
14.6 | -104.6392
14.6037 | -104.6358
14.6074 | -104.6275
14.6111 | -104.6195
14.6148 | -104.6116
14.6185 | -104.6015
14.6222 | -104.5896
14.6259 | -104.5799
14.6296 | -104.5722
14.6333 | -104.5646
14.637 | -104.5594
14.6407 | -104.5563
14.6444 | -104.5559
14.6481 | -104.5592
14.6519 | -104.5611
14.6556 | -104.5648
14.6593 | -104.57
14.663 | -104.5761
14.6667 | -104.5857
14.6704 | -104.5964
14.6741 | -104.6072
14.6778 | -104.6198
14.6815 | -104.6334
14.6852 | -104.6541
14.6889 | -104.6725
14.6926 | -104.6912
14.6963 | -104.7089
14.7 | -104.7255
14.7037 | -104.7425
14.7074 | -104.7577
14.7111 | -104.769
14.7148 | -104.7812
14.7185 | -104.7954
14.7222 | -104.8078
14.7259 | -104.8199
14.7296 | -104.8317
14.7333 | -104.8458
14.737 | -104.8599
14.7407 | -104.8741
14.7444 | -104.886
14.7481 | -104.8974
14.7519 | -104.9095
14.7556 | -104.9196
14.7593 | -104.9232
14.763 | -104.9275
14.7667 | -104.9365
14.7704 | -104.9482
14.7741 | -104.9602
14.7778 | -104.9704
14.7815 | -104.9838
14.7852 | -104.9999
14.7889 | -105.0139
14.7926 | -105.0281
14.7963 | -105.0426
14.8 | -105.0598
14.8037 | -105.0797
14.8074 | -105.0968
14.8111 | -105.1105
14.8148 | -105.1228
14.8185 | -105.1368
14.8222 | -105.149
14.8259 | -105.1622
14.8296 | -105.1761
14.8333 | -105.1901
14.837 | -105.2019
14.8407 | -105.2112
14.8444 | -105.2197
14.8481 | -105.2255
14.8519 | -105.2252
14.8556 | -105.2223
14.8593 | -105.2196
14.863 | -105.2183
14.8667 | -105.2139
14.8704 | -105.2095
14.8741 | -105.2042
14.8778 | -105.1989
14.8815 | -105.1916
14.8852 | -105.1776
14.8889 | -105.1641
14.8926 | -105.1525
14.8963 | -105.1367
14.9 | -105.1203
14.9037 | -105.105
14.9074 | -105.0895
14.9111 | -105.0776
14.9148 | -105.0663
14.9185 | -105.0547
14.9222 | -105.0446
14.9259 | -105.0342
14.9296 | -105.0207
14.9333 | -105.0082
14.937 | -104.9948
14.9407 | -104.9823
14.9444 | -104.9726
14.9481 | -104.9618
14.9519 | -104.9514
14.9556 | -104.9424
14.9593 | -104.942
14.963 | -104.9376
14.9667 | -104.933
14.9704 | -104.9306
14.9741 | -104.9323
14.9778 | -104.9338
14.9815 | -104.9326
14.9852 | -104.9181
14.9889 | -104.8934
14.9926 | -104.886
14.9963 | -104.8792
15 | -104.8708
15.0037 | -104.864
15.0074 | -104.857
15.0111 | -104.8456
15.0148 | -104.832
15.0185 | -104.8168
15.0222 | -104.8031
15.0259 | -104.7872
15.0296 | -104.7706
15.0333 | -104.7503
15.037 | -104.7285
15.0407 | -104.7044
15.0444 | -104.6778
15.0481 | -104.6535
15.0519 | -104.6258
15.0556 | -104.5978
15.0593 | -104.5731
15.063 | -104.5465
15.0667 | -104.5214
15.0704 | -104.4977
15.0741 | -104.4723
15.0778 | -104.4501
15.0815 | -104.4346
15.0852 | -104.4211
15.0889 | -104.4109
15.0926 | -104.4032
15.0963 | -104.3986
15.1 | -104.3996
15.1037 | -104.4027
15.1074 | -104.4084
15.1111 | -104.4205
15.1148 | -104.437
15.1185 | -104.4552
15.1222 | -104.477
15.1259 | -104.5009
15.1296 | -104.5258
15.1333 | -104.5555
15.137 | -104.5858
15.1407 | -104.6168
15.1444 | -104.6512
15.1481 | -104.6894
15.1519 | -104.7297
15.1556 | -104.7686
15.1593 | -104.8058
15.163 | -104.8454
15.1667 | -104.8824
15.1704 | -104.9168
15.1741 | -104.95
15.1778 | -104.979
15.1815 | -105.0074
15.1852 | -105.0367
15.1889 | -105.0633
15.1926 | -105.0842
15.1963 | -105.0994
15.2 | -105.1138
15.2037 | -105.1355
15.2074 | -105.1617
15.2111 | -105.1664
15.2148 | -105.1662
15.2185 | -105.1643
15.2222 | -105.1611
15.2259 | -105.155
15.2296 | -105.1483
15.2333 | -105.1395
15.237 | -105.1302
15.2407 | -105.1173
15.2444 | -105.1016
15.2481 | -105.0863
15.2519 | -105.0717
15.2556 | -105.0559
15.2593 | -105.042
15.263 | -105.0318
15.2667 | -105.0156
15.2704 | -105.0014
15.2741 | -104.9876
15.2778 | -104.9711
15.2815 | -104.9598
15.2852 | -104.9495
15.2889 | -104.9406
15.2926 | -104.9358
15.2963 | -104.9284
15.3 | -104.9194
15.3037 | -104.9138
15.3074 | -104.9112
15.3111 | -104.9121
15.3148 | -104.918
15.3185 | -104.9255
15.3222 | -104.9315
15.3259 | -104.9382
15.3296 | -104.9459
15.3333 | -104.9546
15.337 | -104.9647
15.3407 | -104.9757
15.3444 | -104.9892
15.3481 | -105.0043
15.3519 | -105.0169
15.3556 | -105.0311
15.3593 | -105.0468
15.363 | -105.0612
15.3667 | -105.0714
15.3704 | -105.0799
15.3741 | -105.0871
15.3778 | -105.0932
15.3815 | -105.0999
15.3852 | -105.1091
15.3889 | -105.1186
15.3926 | -105.1272
15.3963 | -105.134
15.4 | -105.1373
15.4037 | -105.1403
15.4074 | -105.1466
15.4111 | -105.154
15.4148 | -105.1638
15.4185 | -105.1707
15.4222 | -105.1773
15.4259 | -105.1813
15.4296 | -105.1824
15.4333 | -105.1792
15.437 | -105.1759
15.4407 | -105.1736
15.4444 | -105.1723
15.4481 | -105.1732
15.4519 | -105.175
15.4556 | -105.1779
15.4593 | -105.181
15.463 | -105.1852
15.4667 | -105.1864
15.4704 | -105.1874
15.4741 | -105.1882
15.4778 | -105.1859
15.4815 | -105.1795
15.4852 | -105.1777
15.4889 | -105.1744
15.4926 | -105.1647
15.4963 | -105.16
15.5 | -105.1556
15.5037 | -105.1483
15.5074 | -105.1421
15.5111 | -105.1336
15.5148 | -105.1232
15.5185 | -105.1167
15.5222 | -105.1138
15.5259 | -105.1085
15.5296 | -105.1004
15.5333 | -105.0908
15.537 | -105.0834
15.5407 | -105.0781
15.5444 | -105.0752
15.5481 | -105.0728
15.5519 | -105.0719
15.5556 | -105.0695
15.5593 | -105.0661
15.563 | -105.0637
15.5667 | -105.0589
15.5704 | -105.0521
15.5741 | -105.0472
15.5778 | -105.0436
15.5815 | -105.0384
15.5852 | -105.0323
15.5889 | -105.0272
15.5926 | -105.022
15.5963 | -105.0179
15.6 | -105.0126
15.6037 | -105.002
15.6074 | -104.9899
15.6111 | -104.9788
15.6148 | -104.9676
15.6185 | -104.9565
15.6222 | -104.9491
15.6259 | -104.9386
15.6296 | -104.924
15.6333 | -104.9105
15.637 | -104.8969
15.6407 | -104.8879
15.6444 | -104.8804
15.6481 | -104.8729
15.6519 | -104.8659
15.6556 | -104.861
15.6593 | -104.8557
15.663 | -104.8506
15.6667 | -104.8452
15.6704 | -104.8395
15.6741 | -104.8358
15.6778 | -104.8305
15.6815 | -104.8259
15.6852 | -104.8238
15.6889 | -104.8247
15.6926 | -104.8278
15.6963 | -104.8304
15.7 | -104.8324
15.7037 | -104.8403
15.7074 | -104.8483
15.7111 | -104.8551
15.7148 | -104.8632
15.7185 | -104.8725
15.7222 | -104.8778
15.7259 | -104.8829
15.7296 | -104.8892
15.7333 | -104.8956
15.737 | -104.9028
15.7407 | -104.911
15.7444 | -104.9182
15.7481 | -104.922
15.7519 | -104.9258
15.7556 | -104.9281
15.7593 | -104.9297
15.763 | -104.9296
15.7667 | -104.9298
15.7704 | -104.928
15.7741 | -104.9233
15.7778 | -104.9155
15.7815 | -104.9107
15.7852 | -104.903
15.7889 | -104.8926
15.7926 | -104.8847
15.7963 | -104.8772
15.8 | -104.8669
15.8037 | -104.8538
15.8074 | -104.8413
15.8111 | -104.8266
15.8148 | -104.8112
15.8185 | -104.7961
15.8222 | -104.7816
15.8259 | -104.7701
15.8296 | -104.7582
15.8333 | -104.7456
15.837 | -104.7345
15.8407 | -104.7258
15.8444 | -104.7182
15.8481 | -104.7128
15.8519 | -104.7104
15.8556 | -104.7113
15.8593 | -104.7114
15.863 | -104.712
15.8667 | -104.7116
15.8704 | -104.7118
15.8741 | -104.7134
15.8778 | -104.7176
15.8815 | -104.7259
15.8852 | -104.7347
15.8889 | -104.7427
15.8926 | -104.7521
15.8963 | -104.7599
15.9 | -104.7659
15.9037 | -104.7746
15.9074 | -104.7823
15.9111 | -104.7878
15.9148 | -104.7938
15.9185 | -104.8008
15.9222 | -104.8094
15.9259 | -104.8189
15.9296 | -104.8257
15.9333 | -104.8307
15.937 | -104.8376
15.9407 | -104.8449
15.9444 | -104.8515
15.9481 | -104.8616
15.9519 | -104.8714
15.9556 | -104.8796
15.9593 | -104.8839
15.963 | -104.8904
15.9667 | -104.8951
15.9704 | -104.8991
15.9741 | -104.9011
15.9778 | -104.9004
15.9815 | -104.901
15.9852 | -104.9018
15.9889 | -104.8991
15.9926 | -104.8926
15.9963 | -104.8866
16 | -104.8841
16.0037 | -104.8808
16.0074 | -104.8734
16.0111 | -104.865
16.0148 | -104.8566
16.0185 | -104.8445
16.0222 | -104.8298
16.0259 | -104.8142
16.0296 | -104.7939
16.0333 | -104.7793
16.037 | -104.765
16.0407 | -104.7519
16.0444 | -104.738
16.0481 | -104.7249
16.0519 | -104.7124
16.0556 | -104.6975
16.0593 | -104.6851
16.063 | -104.6742
16.0667 | -104.6637
16.0704 | -104.6544
16.0741 | -104.6473
16.0778 | -104.6415
16.0815 | -104.6371
16.0852 | -104.6338
16.0889 | -104.6348
16.0926 | -104.6393
16.0963 | -104.643
16.1 | -104.6487
16.1037 | -104.66
16.1074 | -104.6735
16.1111 | -104.6862
16.1148 | -104.6996
16.1185 | -104.718
16.1222 | -104.7355
16.1259 | -104.7529
16.1296 | -104.7739
16.1333 | -104.7944
16.137 | -104.817
16.1407 | -104.8406
16.1444 | -104.8635
16.1481 | -104.8835
16.1519 | -104.9013
16.1556 | -104.9197
16.1593 | -104.9374
16.163 | -104.9534
16.1667 | -104.9687
16.1704 | -104.9804
16.1741 | -104.9904
16.1778 | -104.9971
16.1815 | -105.0023
16.1852 | -105.0057
16.1889 | -105.0111
16.1926 | -105.0129
16.1963 | -105.0128
16.2 | -105.0117
16.2037 | -105.0123
16.2074 | -105.0126
16.2111 | -105.0078
16.2148 | -104.9994
16.2185 | -104.9898
16.2222 | -104.9805
16.2259 | -104.9722
16.2296 | -104.9603
16.2333 | -104.9505
16.237 | -104.9403
16.2407 | -104.9324
16.2444 | -104.9228
16.2481 | -104.9129
16.2519 | -104.9075
16.2556 | -104.9038
16.2593 | -104.8992
16.263 | -104.8927
16.2667 | -104.8871
16.2704 | -104.8832
16.2741 | -104.8805
16.2778 | -104.8774
16.2815 | -104.8756
16.2852 | -104.8749
16.2889 | -104.8747
16.2926 | -104.8761
16.2963 | -104.875
16.3 | -104.874
16.3037 | -104.8739
16.3074 | -104.8739
16.3111 | -104.8745
16.3148 | -104.8777
16.3185 | -104.8837
16.3222 | -104.8888
16.3259 | -104.8915
16.3296 | -104.8948
16.3333 | -104.8956
16.337 | -104.8893
16.3407 | -104.8846
16.3444 | -104.8843
16.3481 | -104.8847
16.3519 | -104.8814
16.3556 | -104.8771
16.3593 | -104.8752
16.363 | -104.872
16.3667 | -104.8693
16.3704 | -104.8662
16.3741 | -104.8624
16.3778 | -104.857
16.3815 | -104.8494
16.3852 | -104.8436
16.3889 | -104.8409
16.3926 | -104.8372
16.3963 | -104.8353
16.4 | -104.8302
16.4037 | -104.8242
16.4074 | -104.8186
16.4111 | -104.8171
16.4148 | -104.8169
16.4185 | -104.8155
16.4222 | -104.8144
16.4259 | -104.815
16.4296 | -104.8146
16.4333 | -104.8121
16.437 | -104.8105
16.4407 | -104.8116
16.4444 | -104.8156
16.4481 | -104.8192
16.4519 | -104.8239
16.4556 | -104.828
16.4593 | -104.8306
16.463 | -104.8331
16.4667 | -104.8367
16.4704 | -104.8439
16.4741 | -104.8481
16.4778 | -104.8481
16.4815 | -104.8473
16.4852 | -104.8468
16.4889 | -104.8454
16.4926 | -104.8443
16.4963 | -104.8429
16.5 | -104.8401
16.5037 | -104.8368
16.5074 | -104.8338
16.5111 | -104.8318
16.5148 | -104.8274
16.5185 | -104.8213
16.5222 | -104.8151
16.5259 | -104.8074
16.5296 | -104.8023
16.5333 | -104.7941
16.537 | -104.7827
16.5407 | -104.7702
16.5444 | -104.7588
16.5481 | -104.7467
16.5519 | -104.7359
16.5556 | -104.7277
16.5593 | -104.7205
16.563 | -104.7157
16.5667 | -104.7097
16.5704 | -104.7063
16.5741 | -104.7036
16.5778 | -104.6983
16.5815 | -104.6895
16.5852 | -104.682
16.5889 | -104.6777
16.5926 | -104.6766
16.5963 | -104.6744
16.6 | -104.673
16.6037 | -104.67
16.6074 | -104.6675
16.6111 | -104.6652
16.6148 | -104.6631
16.6185 | -104.6669
16.6222 | -104.6724
16.6259 | -104.6757
16.6296 | -104.6806
16.6333 | -104.6856
16.637 | -104.6933
16.6407 | -104.7063
16.6444 | -104.7199
16.6481 | -104.7305
16.6519 | -104.7394
16.6556 | -104.7521
16.6593 | -104.7673
16.663 | -104.782
16.6667 | -104.7997
16.6704 | -104.8184
16.6741 | -104.8355
16.6778 | -104.8514
16.6815 | -104.8656
16.6852 | -104.8794
16.6889 | -104.8935
16.6926 | -104.9084
16.6963 | -104.9225
16.7 | -104.9372
16.7037 | -104.951
16.7074 | -104.9583
16.7111 | -104.9663
16.7148 | -104.982
16.7185 | -104.9973
16.7222 | -105.0073
16.7259 | -105.0171
16.7296 | -105.0272
16.7333 | -105.0358
16.737 | -105.0412
16.7407 | -105.0441
16.7444 | -105.0447
16.7481 | -105.0413
16.7519 | -105.0384
16.7556 | -105.0362
16.7593 | -105.0327
16.763 | -105.0288
16.7667 | -105.0209
16.7704 | -105.0093
16.7741 | -104.9926
16.7778 | -104.9798
16.7815 | -104.9651
16.7852 | -104.948
16.7889 | -104.9298
16.7926 | -104.9107
16.7963 | -104.8921
16.8 | -104.8746
16.8037 | -104.8568
16.8074 | -104.8393
16.8111 | -104.8231
16.8148 | -104.8079
16.8185 | -104.7936
16.8222 | -104.7811
16.8259 | -104.7695
16.8296 | -104.7578
16.8333 | -104.7493
16.837 | -104.7415
16.8407 | -104.7326
16.8444 | -104.7268
16.8481 | -104.7241
16.8519 | -104.7242
16.8556 | -104.7283
16.8593 | -104.7314
16.863 | -104.7333
16.8667 | -104.7358
16.8704 | -104.7432
16.8741 | -104.7527
16.8778 | -104.7642
16.8815 | -104.7765
16.8852 | -104.7901
16.8889 | -104.8052
16.8926 | -104.8255
16.8963 | -104.8455
16.9 | -104.8617
16.9037 | -104.8792
16.9074 | -104.9001
16.9111 | -104.9245
16.9148 | -104.9495
16.9185 | -104.9763
16.9222 | -105.0057
16.9259 | -105.0367
16.9296 | -105.0647
16.9333 | -105.0917
16.937 | -105.1192
16.9407 | -105.1456
16.9444 | -105.171
16.9481 | -105.2021
16.9519 | -105.2335
16.9556 | -105.2579
16.9593 | -105.2798
16.963 | -105.3034
16.9667 | -105.3267
16.9704 | -105.3491
16.9741 | -105.3711
16.9778 | -105.3918
16.9815 | -105.4106
16.9852 | -105.4308
16.9889 | -105.4531
16.9926 | -105.4706
16.9963 | -105.4852
17 | -105.5013
17.0037 | -105.5145
17.0074 | -105.525
17.0111 | -105.5368
17.0148 | -105.5531
17.0185 | -105.559
17.0222 | -105.5584
17.0259 | -105.5585
17.0296 | -105.5578
17.0333 | -105.5524
17.037 | -105.5461
17.0407 | -105.5396
17.0444 | -105.533
17.0481 | -105.5255
17.0519 | -105.5153
17.0556 | -105.5038
17.0593 | -105.4907
17.063 | -105.4744
17.0667 | -105.4579
17.0704 | -105.4388
17.0741 | -105.4165
17.0778 | -105.3944
17.0815 | -105.3744
17.0852 | -105.3539
17.0889 | -105.3296
17.0926 | -105.3031
17.0963 | -105.2811
17.1 | -105.2575
17.1037 | -105.2334
17.1074 | -105.2129
17.1111 | -105.1926
17.1148 | -105.1678
17.1185 | -105.1444
17.1222 | -105.1262
17.1259 | -105.1094
17.1296 | -105.0945
17.1333 | -105.0816
17.137 | -105.0702
17.1407 | -105.062
17.1444 | -105.0597
17.1481 | -105.0595
17.1519 | -105.0569
17.1556 | -105.0541
17.1593 | -105.0548
17.163 | -105.0559
17.1667 | -105.0576
17.1704 | -105.0585
17.1741 | -105.0619
17.1778 | -105.0677
17.1815 | -105.0746
17.1852 | -105.0817
17.1889 | -105.0888
17.1926 | -105.0967
17.1963 | -105.1051
17.2 | -105.113
17.2037 | -105.1245
17.2074 | -105.1353
17.2111 | -105.1434
17.2148 | -105.1523
17.2185 | -105.1607
17.2222 | -105.1714
17.2259 | -105.1796
17.2296 | -105.1821
17.2333 | -105.1804
17.237 | -105.1755
17.2407 | -105.1685
17.2444 | -105.1616
17.2481 | -105.1581
17.2519 | -105.1585
17.2556 | -105.1572
17.2593 | -105.1492
17.263 | -105.1401
17.2667 | -105.1328
17.2704 | -105.1236
17.2741 | -105.1093
17.2778 | -105.0926
17.2815 | -105.0723
17.2852 | -105.0455
17.2889 | -105.0201
17.2926 | -104.9982
17.2963 | -104.9754
17.3 | -104.9533
17.3037 | -104.931
17.3074 | -104.9071
17.3111 | -104.8808
17.3148 | -104.8571
17.3185 | -104.8333
17.3222 | -104.8108
17.3259 | -104.7874
17.3296 | -104.7626
17.3333 | -104.7386
17.337 | -104.7161
17.3407 | -104.6884
17.3444 | -104.6598
17.3481 | -104.6348
17.3519 | -104.6134
17.3556 | -104.5901
17.3593 | -104.5689
17.363 | -104.5523
17.3667 | -104.534
17.3704 | -104.5147
17.3741 | -104.4937
17.3778 | -104.4636
17.3815 | -104.4324
17.3852 | -104.4084
17.3889 | -104.3855
17.3926 | -104.3613
17.3963 | -104.3414
17.4 | -104.3258
17.4037 | -104.3107
17.4074 | -104.2945
17.4111 | -104.2798
17.4148 | -104.2669
17.4185 | -104.2546
17.4222 | -104.2438
17.4259 | -104.2301
17.4296 | -104.2141
17.4333 | -104.2003
17.437 | -104.1865
17.4407 | -104.1706
17.4444 | -104.153
17.4481 | -104.1366
17.4519 | -104.1213
17.4556 | -104.1068
17.4593 | -104.0944
17.463 | -104.081
17.4667 | -104.068
17.4704 | -104.0552
17.4741 | -104.0461
17.4778 | -104.038
17.4815 | -104.0254
17.4852 | -104.0148
17.4889 | -104.0109
17.4926 | -104.0069
17.4963 | -104.0029
17.5 | -103.9996
17.5037 | -103.9985
17.5074 | -104.0018
17.5111 | -104.0071
17.5148 | -104.0113
17.5185 | -104.0162
17.5222 | -104.0249
17.5259 | -104.0383
17.5296 | -104.0552
17.5333 | -104.0741
17.537 | -104.0899
17.5407 | -104.1007
17.5444 | -104.1109
17.5481 | -104.122
17.5519 | -104.1337
17.5556 | -104.1496
17.5593 | -104.1677
17.563 | -104.1837
17.5667 | -104.2009
17.5704 | -104.2174
17.5741 | -104.2321
17.5778 | -104.2455
17.5815 | -104.2603
17.5852 | -104.2751
17.5889 | -104.2938
17.5926 | -104.315
17.5963 | -104.3334
17.6 | -104.3498
17.6037 | -104.3696
17.6074 | -104.3913
17.6111 | -104.4094
17.6148 | -104.4277
17.6185 | -104.4486
17.6222 | -104.4725
17.6259 | -104.5049
17.6296 | -104.5254
17.6333 | -104.557
17.637 | -104.5799
17.6407 | -104.6009
17.6444 | -104.6203
17.6481 | -104.6366
17.6519 | -104.6537
17.6556 | -104.6704
17.6593 | -104.6887
17.663 | -104.7086
17.6667 | -104.7277
17.6704 | -104.7475
17.6741 | -104.7651
17.6778 | -104.7822
17.6815 | -104.7983
17.6852 | -104.8113
17.6889 | -104.8254
17.6926 | -104.8416
17.6963 | -104.8591
17.7 | -104.8732
17.7037 | -104.8822
17.7074 | -104.8898
17.7111 | -104.8973
17.7148 | -104.9027
17.7185 | -104.9049
17.7222 | -104.9091
17.7259 | -104.9131
17.7296 | -104.9179
17.7333 | -104.9255
17.737 | -104.9314
17.7407 | -104.927
17.7444 | -104.9185
17.7481 | -104.9076
17.7519 | -104.8975
17.7556 | -104.888
17.7593 | -104.8793
17.763 | -104.871
17.7667 | -104.862
17.7704 | -104.8532
17.7741 | -104.8429
17.7778 | -104.8333
17.7815 | -104.821
17.7852 | -104.8083
17.7889 | -104.7957
17.7926 | -104.7847
17.7963 | -104.7754
17.8 | -104.7656
17.8037 | -104.7528
17.8074 | -104.7385
17.8111 | -104.726
17.8148 | -104.7178
17.8185 | -104.7089
17.8222 | -104.702
17.8259 | -104.7002
17.8296 | -104.7024
17.8333 | -104.7041
17.837 | -104.704
17.8407 | -104.7023
17.8444 | -104.6996
17.8481 | -104.6991
17.8519 | -104.699
17.8556 | -104.6947
17.8593 | -104.6904
17.863 | -104.6934
17.8667 | -104.6995
17.8704 | -104.6972
17.8741 | -104.6909
17.8778 | -104.6893
17.8815 | -104.691
17.8852 | -104.6921
17.8889 | -104.6916
17.8926 | -104.6882
17.8963 | -104.6861
17.9 | -104.685
17.9037 | -104.6804
17.9074 | -104.6734
17.9111 | -104.6663
17.9148 | -104.6547
17.9185 | -104.6425
17.9222 | -104.6301
17.9259 | -104.617
17.9296 | -104.6066
17.9333 | -104.5995
17.937 | -104.5926
17.9407 | -104.5838
17.9444 | -104.5738
17.9481 | -104.5622
17.9519 | -104.5501
17.9556 | -104.5415
17.9593 | -104.5347
17.963 | -104.5293
17.9667 | -104.523
17.9704 | -104.5204
17.9741 | -104.5218
17.9778 | -104.5205
17.9815 | -104.5172
17.9852 | -104.5149
17.9889 | -104.5176
17.9926 | -104.5221
17.9963 | -104.5265
18 | -104.5292
18.0037 | -104.5312
18.0074 | -104.5327
18.0111 | -104.5386
18.0148 | -104.5455
18.0185 | -104.5548
18.0222 | -104.5668
18.0259 | -104.5802
18.0296 | -104.5933
18.0333 | -104.6035
18.037 | -104.6178
18.0407 | -104.6335
18.0444 | -104.6493
18.0481 | -104.6665
18.0519 | -104.682
18.0556 | -104.6924
18.0593 | -104.7026
18.063 | -104.714
18.0667 | -104.7221
18.0704 | -104.7269
18.0741 | -104.7309
18.0778 | -104.7345
18.0815 | -104.7344
18.0852 | -104.7312
18.0889 | -104.7303
18.0926 | -104.7308
18.0963 | -104.7312
18.1 | -104.7283
18.1037 | -104.7256
18.1074 | -104.7235
18.1111 | -104.7222
18.1148 | -104.7161
18.1185 | -104.7062
18.1222 | -104.6984
18.1259 | -104.6939
18.1296 | -104.69
18.1333 | -104.6815
18.137 | -104.6718
18.1407 | -104.6614
18.1444 | -104.651
18.1481 | -104.6445
18.1519 | -104.6411
18.1556 | -104.6373
18.1593 | -104.635
18.163 | -104.6342
18.1667 | -104.6377
18.1704 | -104.6365
18.1741 | -104.6381
18.1778 | -104.6402
18.1815 | -104.6399
18.1852 | -104.6358
18.1889 | -104.6322
18.1926 | -104.6316
18.1963 | -104.6331
18.2 | -104.6365
18.2037 | -104.6382
18.2074 | -104.6418
18.2111 | -104.6464
18.2148 | -104.6483
18.2185 | -104.6502
18.2222 | -104.6516
18.2259 | -104.6551
18.2296 | -104.6557
18.2333 | -104.6582
18.237 | -104.6622
18.2407 | -104.6661
18.2444 | -104.6697
18.2481 | -104.6719
18.2519 | -104.6744
18.2556 | -104.6805
18.2593 | -104.6894
18.263 | -104.6982
18.2667 | -104.7063
18.2704 | -104.7101
18.2741 | -104.7127
18.2778 | -104.7147
18.2815 | -104.7182
18.2852 | -104.7225
18.2889 | -104.728
18.2926 | -104.736
18.2963 | -104.7422
18.3 | -104.7499
18.3037 | -104.7592
18.3074 | -104.768
18.3111 | -104.7771
18.3148 | -104.79
18.3185 | -104.8046
18.3222 | -104.812
18.3259 | -104.8182
18.3296 | -104.828
18.3333 | -104.839
18.337 | -104.8491
18.3407 | -104.8595
18.3444 | -104.8712
18.3481 | -104.8787
18.3519 | -104.8823
18.3556 | -104.8851
18.3593 | -104.8917
18.363 | -104.8961
18.3667 | -104.8972
18.3704 | -104.9001
18.3741 | -104.9063
18.3778 | -104.9116
18.3815 | -104.9131
18.3852 | -104.9125
18.3889 | -104.9108
18.3926 | -104.9088
18.3963 | -104.9051
18.4 | -104.899
18.4037 | -104.892
18.4074 | -104.885
18.4111 | -104.8757
18.4148 | -104.8671
18.4185 | -104.8579
18.4222 | -104.8473
18.4259 | -104.8357
18.4296 | -104.8264
18.4333 | -104.8182
18.437 | -104.8086
18.4407 | -104.7944
18.4444 | -104.7849
18.4481 | -104.7794
18.4519 | -104.771
18.4556 | -104.7564
18.4593 | -104.7412
18.463 | -104.7281
18.4667 | -104.7136
18.4704 | -104.6965
18.4741 | -104.6851
18.4778 | -104.6776
18.4815 | -104.6663
18.4852 | -104.6577
18.4889 | -104.6522
18.4926 | -104.6507
18.4963 | -104.6465
18.5 | -104.6384
18.5037 | -104.6288
18.5074 | -104.6218
18.5111 | -104.6134
18.5148 | -104.6004
18.5185 | -104.588
18.5222 | -104.5808
18.5259 | -104.5742
18.5296 | -104.5659
18.5333 | -104.5545
18.537 | -104.5434
18.5407 | -104.5313
18.5444 | -104.5169
18.5481 | -104.5037
18.5519 | -104.4896
18.5556 | -104.4712
18.5593 | -104.4505
18.563 | -104.4319
18.5667 | -104.4128
18.5704 | -104.3925
18.5741 | -104.3721
18.5778 | -104.3505
18.5815 | -104.3263
18.5852 | -104.3043
18.5889 | -104.2843
18.5926 | -104.2642
18.5963 | -104.2423
18.6 | -104.2181
18.6037 | -104.1946
18.6074 | -104.1709
18.6111 | -104.1511
18.6148 | -104.1352
18.6185 | -104.1212
18.6222 | -104.1075
18.6259 | -104.0945
18.6296 | -104.0836
18.6333 | -104.0728
18.637 | -104.0642
18.6407 | -104.0571
18.6444 | -104.0516
18.6481 | -104.0501
18.6519 | -104.0507
18.6556 | -104.0486
18.6593 | -104.0497
18.663 | -104.0549
18.6667 | -104.0614
18.6704 | -104.0685
18.6741 | -104.0768
18.6778 | -104.0868
18.6815 | -104.096
18.6852 | -104.1025
18.6889 | -104.1052
18.6926 | -104.1054
18.6963 | -104.1058
18.7 | -104.0955
18.7037 | -104.0882
18.7074 | -104.0876
18.7111 | -104.0887
18.7148 | -104.0888
18.7185 | -104.0869
18.7222 | -104.082
18.7259 | -104.0742
18.7296 | -104.0672
18.7333 | -104.0603
18.737 | -104.0477
18.7407 | -104.0283
18.7444 | -104.009
18.7481 | -103.9915
18.7519 | -103.9709
18.7556 | -103.9448
18.7593 | -103.9171
18.763 | -103.8901
18.7667 | -103.8607
18.7704 | -103.8224
18.7741 | -103.7786
18.7778 | -103.7402
18.7815 | -103.7057
18.7852 | -103.6681
18.7889 | -103.6259
18.7926 | -103.5864
18.7963 | -103.5494
18.8 | -103.515
18.8037 | -103.4807
18.8074 | -103.4469
18.8111 | -103.413
18.8148 | -103.3763
18.8185 | -103.3421
18.8222 | -103.3133
18.8259 | -103.2868
18.8296 | -103.2595
18.8333 | -103.2337
18.837 | -103.2118
18.8407 | -103.1937
18.8444 | -103.1798
18.8481 | -103.1698
18.8519 | -103.1615
18.8556 | -103.152
18.8593 | -103.1432
18.863 | -103.1401
18.8667 | -103.142
18.8704 | -103.144
18.8741 | -103.1435
18.8778 | -103.1432
18.8815 | -103.148
18.8852 | -103.1557
18.8889 | -103.1637
18.8926 | -103.1747
18.8963 | -103.1861
18.9 | -103.1956
18.9037 | -103.2058
18.9074 | -103.2198
18.9111 | -103.2356
18.9148 | -103.2517
18.9185 | -103.2676
18.9222 | -103.2857
18.9259 | -103.3042
18.9296 | -103.3224
18.9333 | -103.341
18.937 | -103.3582
18.9407 | -103.3748
18.9444 | -103.3921
18.9481 | -103.4118
18.9519 | -103.4355
18.9556 | -103.4618
18.9593 | -103.4856
18.963 | -103.5055
18.9667 | -103.5235
18.9704 | -103.5463
18.9741 | -103.5717
18.9778 | -103.5893
18.9815 | -103.5997
18.9852 | -103.6099
18.9889 | -103.6234
18.9926 | -103.6424
18.9963 | -103.6625
19 | -103.6779
19.0037 | -103.6929
19.0074 | -103.7127
19.0111 | -103.7332
19.0148 | -103.7514
19.0185 | -103.7693
19.0222 | -103.7876
19.0259 | -103.8073
19.0296 | -103.8265
19.0333 | -103.8423
19.037 | -103.8614
19.0407 | -103.8898
19.0444 | -103.9189
19.0481 | -103.9379
19.0519 | -103.9512
19.0556 | -103.9636
19.0593 | -103.9769
19.063 | -103.9884
19.0667 | -103.999
19.0704 | -104.0097
19.0741 | -104.0213
19.0778 | -104.0345
19.0815 | -104.0489
19.0852 | -104.0655
19.0889 | -104.0825
19.0926 | -104.0944
19.0963 | -104.1056
19.1 | -104.115
19.1037 | -104.1243
19.1074 | -104.1334
19.1111 | -104.1417
19.1148 | -104.1482
19.1185 | -104.1538
19.1222 | -104.1626
19.1259 | -104.1733
19.1296 | -104.1816
19.1333 | -104.1871
19.137 | -104.1948
19.1407 | -104.2067
19.1444 | -104.221
19.1481 | -104.2303
19.1519 | -104.2341
19.1556 | -104.2388
19.1593 | -104.2462
19.163 | -104.2541
19.1667 | -104.2659
19.1704 | -104.2755
19.1741 | -104.2801
19.1778 | -104.2842
19.1815 | -104.2903
19.1852 | -104.2951
19.1889 | -104.3023
19.1926 | -104.3107
19.1963 | -104.3143
19.2 | -104.3127
19.2037 | -104.3061
19.2074 | -104.302
19.2111 | -104.3023
19.2148 | -104.3042
19.2185 | -104.301
19.2222 | -104.2928
19.2259 | -104.2848
19.2296 | -104.2814
19.2333 | -104.2714
19.237 | -104.2668
19.2407 | -104.2609
19.2444 | -104.2523
19.2481 | -104.243
19.2519 | -104.2314
19.2556 | -104.222
19.2593 | -104.2143
19.263 | -104.204
19.2667 | -104.1928
19.2704 | -104.1821
19.2741 | -104.1726
19.2778 | -104.1642
19.2815 | -104.1549
19.2852 | -104.1468
19.2889 | -104.1406
19.2926 | -104.1343
19.2963 | -104.1286
19.3 | -104.123
19.3037 | -104.1204
19.3074 | -104.1228
19.3111 | -104.1267
19.3148 | -104.1275
19.3185 | -104.1247
19.3222 | -104.1204
19.3259 | -104.1167
19.3296 | -104.1107
19.3333 | -104.0997
19.337 | -104.0864
19.3407 | -104.0749
19.3444 | -104.0623
19.3481 | -104.0487
19.3519 | -104.034
19.3556 | -104.0187
19.3593 | -104.0002
19.363 | -103.979
19.3667 | -103.9595
19.3704 | -103.9346
19.3741 | -103.9086
19.3778 | -103.8804
19.3815 | -103.8467
19.3852 | -103.8081
19.3889 | -103.7678
19.3926 | -103.7255
19.3963 | -103.6785
19.4 | -103.6323
19.4037 | -103.5867
19.4074 | -103.5369
19.4111 | -103.4804
19.4148 | -103.4242
19.4185 | -103.3698
19.4222 | -103.3141
19.4259 | -103.2573
19.4296 | -103.2021
19.4333 | -103.1472
19.437 | -103.092
19.4407 | -103.0369
19.4444 | -102.9894
19.4481 | -102.9449
19.4519 | -102.8975
19.4556 | -102.8498
19.4593 | -102.8075
19.463 | -102.7685
19.4667 | -102.731
19.4704 | -102.6991
19.4741 | -102.6732
19.4778 | -102.6541
19.4815 | -102.6377
19.4852 | -102.6233
19.4889 | -102.6126
19.4926 | -102.6084
19.4963 | -102.6096
19.5 | -102.6132
19.5037 | -102.6192
19.5074 | -102.6304
19.5111 | -102.6446
19.5148 | -102.664
19.5185 | -102.6878
19.5222 | -102.7148
19.5259 | -102.7441
19.5296 | -102.7793
19.5333 | -102.816
19.537 | -102.8506
19.5407 | -102.8868
19.5444 | -102.9251
19.5481 | -102.9645
19.5519 | -103.0075
19.5556 | -103.0529
19.5593 | -103.099
19.563 | -103.1441
19.5667 | -103.1887
19.5704 | -103.2347
19.5741 | -103.2826
19.5778 | -103.3325
19.5815 | -103.3811
19.5852 | -103.428
19.5889 | -103.4727
19.5926 | -103.5194
19.5963 | -103.5698
19.6 | -103.6193
19.6037 | -103.6672
19.6074 | -103.7164
19.6111 | -103.7661
19.6148 | -103.8127
19.6185 | -103.8533
19.6222 | -103.8948
19.6259 | -103.9337
19.6296 | -103.9684
19.6333 | -103.9966
19.637 | -104.0205
19.6407 | -104.0426
19.6444 | -104.062
19.6481 | -104.0817
19.6519 | -104.1017
19.6556 | -104.1168
19.6593 | -104.1281
19.663 | -104.1373
19.6667 | -104.142
19.6704 | -104.1447
19.6741 | -104.1477
19.6778 | -104.1462
19.6815 | -104.1374
19.6852 | -104.1257
19.6889 | -104.1181
19.6926 | -104.1093
19.6963 | -104.0968
19.7 | -104.0843
19.7037 | -104.0711
19.7074 | -104.0516
19.7111 | -104.0277
19.7148 | -104.0056
19.7185 | -103.9824
19.7222 | -103.9551
19.7259 | -103.9264
19.7296 | -103.9015
19.7333 | -103.8793
19.737 | -103.8545
19.7407 | -103.8291
19.7444 | -103.8046
19.7481 | -103.7779
19.7519 | -103.7522
19.7556 | -103.7283
19.7593 | -103.704
19.763 | -103.6792
19.7667 | -103.674
19.7704 | -103.656
19.7741 | -103.642
19.7778 | -103.6298
19.7815 | -103.6223
19.7852 | -103.6179
19.7889 | -103.6131
19.7926 | -103.6106
19.7963 | -103.6125
19.8 | -103.6175
19.8037 | -103.6253
19.8074 | -103.6366
19.8111 | -103.6517
19.8148 | -103.6719
19.8185 | -103.696
19.8222 | -103.7186
19.8259 | -103.7416
19.8296 | -103.7703
19.8333 | -103.7995
19.837 | -103.8269
19.8407 | -103.8576
19.8444 | -103.8914
19.8481 | -103.9257
19.8519 | -103.9624
19.8556 | -103.9991
19.8593 | -104.0339
19.863 | -104.0685
19.8667 | -104.1032
19.8704 | -104.1395
19.8741 | -104.1753
19.8778 | -104.2088
19.8815 | -104.2426
19.8852 | -104.2757
19.8889 | -104.3048
19.8926 | -104.3322
19.8963 | -104.3605
19.9 | -104.389
19.9037 | -104.4114
19.9074 | -104.4319
19.9111 | -104.4539
19.9148 | -104.4764
19.9185 | -104.5006
19.9222 | -104.5217
19.9259 | -104.5397
19.9296 | -104.5553
19.9333 | -104.5649
19.937 | -104.5765
19.9407 | -104.59
19.9444 | -104.6017
19.9481 | -104.6112
19.9519 | -104.62
19.9556 | -104.6259
19.9593 | -104.6302
19.963 | -104.6384
19.9667 | -104.6483
19.9704 | -104.6556
19.9741 | -104.6563
19.9778 | -104.6561
19.9815 | -104.6573
19.9852 | -104.6565
19.9889 | -104.6566
19.9926 | -104.6637
19.9963 | -104.6688
20 | -104.6743
20.0037 | -104.68
20.0074 | -104.6885
20.0111 | -104.6959
20.0148 | -104.7011
20.0185 | -104.7018
20.0222 | -104.7009
20.0259 | -104.7042
20.0296 | -104.7085
20.0333 | -104.7056
20.037 | -104.6997
20.0407 | -104.6948
20.0444 | -104.6916
20.0481 | -104.6857
20.0519 | -104.6792
20.0556 | -104.673
20.0593 | -104.6638
20.063 | -104.6511
20.0667 | -104.6348
20.0704 | -104.6168
20.0741 | -104.6007
20.0778 | -104.5868
20.0815 | -104.5692
20.0852 | -104.549
20.0889 | -104.5276
20.0926 | -104.4996
20.0963 | -104.4672
20.1 | -104.438
20.1037 | -104.4139
20.1074 | -104.3909
20.1111 | -104.3664
20.1148 | -104.3364
20.1185 | -104.3079
20.1222 | -104.2822
20.1259 | -104.255
20.1296 | -104.2262
20.1333 | -104.1986
20.137 | -104.1693
20.1407 | -104.1403
20.1444 | -104.1138
20.1481 | -104.0899
20.1519 | -104.0666
20.1556 | -104.0421
20.1593 | -104.0161
20.163 | -103.9939
20.1667 | -103.9757
20.1704 | -103.9601
20.1741 | -103.9459
20.1778 | -103.9335
20.1815 | -103.9216
20.1852 | -103.9111
20.1889 | -103.904
20.1926 | -103.8996
20.1963 | -103.8964
20.2 | -103.8992
20.2037 | -103.903
20.2074 | -103.9084
20.2111 | -103.9148
20.2148 | -103.9235
20.2185 | -103.9362
20.2222 | -103.9491
20.2259 | -103.9631
20.2296 | -103.9776
20.2333 | -103.9956
20.237 | -104.0159
20.2407 | -104.0369
20.2444 | -104.0591
20.2481 | -104.0836
20.2519 | -104.1103
20.2556 | -104.1338
20.2593 | -104.1594
20.263 | -104.1893
20.2667 | -104.2173
20.2704 | -104.2419
20.2741 | -104.2626
20.2778 | -104.2817
20.2815 | -104.3017
20.2852 | -104.3209
20.2889 | -104.3408
20.2926 | -104.3572
20.2963 | -104.3709
20.3 | -104.3822
20.3037 | -104.3817
20.3074 | -104.3951
20.3111 | -104.4068
20.3148 | -104.414
20.3185 | -104.4218
20.3222 | -104.4314
20.3259 | -104.4353
20.3296 | -104.4358
20.3333 | -104.4336
20.337 | -104.4295
20.3407 | -104.4235
20.3444 | -104.4145
20.3481 | -104.406
20.3519 | -104.3978
20.3556 | -104.3892
20.3593 | -104.3793
20.363 | -104.3687
20.3667 | -104.3554
20.3704 | -104.3436
20.3741 | -104.3319
20.3778 | -104.3213
20.3815 | -104.3127
20.3852 | -104.3055
20.3889 | -104.2953
20.3926 | -104.2806
20.3963 | -104.2682
20.4 | -104.257
20.4037 | -104.2474
20.4074 | -104.237
20.4111 | -104.2278
20.4148 | -104.2204
20.4185 | -104.2151
20.4222 | -104.2117
20.4259 | -104.2095
20.4296 | -104.2062
20.4333 | -104.2027
20.437 | -104.1999
20.4407 | -104.2003
20.4444 | -104.2018
20.4481 | -104.2038
20.4519 | -104.2037
20.4556 | -104.2023
20.4593 | -104.1998
20.463 | -104.1982
20.4667 | -104.1987
20.4704 | -104.2006
20.4741 | -104.2018
20.4778 | -104.2009
20.4815 | -104.1972
20.4852 | -104.1952
20.4889 | -104.1954
20.4926 | -104.194
20.4963 | -104.187
20.5 | -104.1812
20.5037 | -104.1763
20.5074 | -104.1648
20.5111 | -104.1542
20.5148 | -104.1484
20.5185 | -104.1465
20.5222 | -104.142
20.5259 | -104.1344
20.5296 | -104.1258
20.5333 | -104.1173
20.537 | -104.107
20.5407 | -104.0968
20.5444 | -104.0874
20.5481 | -104.076
20.5519 | -104.0617
20.5556 | -104.0498
20.5593 | -104.0392
20.563 | -104.0287
20.5667 | -104.015
20.5704 | -104.0009
20.5741 | -103.9879
20.5778 | -103.9728
20.5815 | -103.9588
20.5852 | -103.9461
20.5889 | -103.9345
20.5926 | -103.9219
20.5963 | -103.9091
20.6 | -103.8968
20.6037 | -103.8853
20.6074 | -103.877
20.6111 | -103.8693
20.6148 | -103.8642
20.6185 | -103.8635
20.6222 | -103.8636
20.6259 | -103.8653
20.6296 | -103.8693
20.6333 | -103.8768
20.637 | -103.885
20.6407 | -103.8962
20.6444 | -103.9115
20.6481 | -103.9283
20.6519 | -103.946
20.6556 | -103.9657
20.6593 | -103.9876
20.663 | -104.0093
20.6667 | -104.0312
20.6704 | -104.0571
20.6741 | -104.0858
20.6778 | -104.1141
20.6815 | -104.1435
20.6852 | -104.1733
20.6889 | -104.2022
20.6926 | -104.2315
20.6963 | -104.2623
20.7 | -104.2914
20.7037 | -104.3174
20.7074 | -104.3424
20.7111 | -104.3675
20.7148 | -104.39
20.7185 | -104.4113
20.7222 | -104.4329
20.7259 | -104.4553
20.7296 | -104.4735
20.7333 | -104.4879
20.737 | -104.5032
20.7407 | -104.5169
20.7444 | -104.5258
20.7481 | -104.5298
20.7519 | -104.5315
20.7556 | -104.5326
20.7593 | -104.5308
20.763 | -104.5271
20.7667 | -104.5212
20.7704 | -104.5176
20.7741 | -104.5147
20.7778 | -104.5076
20.7815 | -104.4963
20.7852 | -104.4855
20.7889 | -104.4775
20.7926 | -104.4672
20.7963 | -104.4544
20.8 | -104.4448
20.8037 | -104.4309
20.8074 | -104.4115
20.8111 | -104.3929
20.8148 | -104.374
20.8185 | -104.3551
20.8222 | -104.3392
20.8259 | -104.3244
20.8296 | -104.3099
20.8333 | -104.2943
20.837 | -104.2828
20.8407 | -104.2668
20.8444 | -104.2532
20.8481 | -104.2395
20.8519 | -104.2226
20.8556 | -104.2062
20.8593 | -104.1912
20.863 | -104.1767
20.8667 | -104.1616
20.8704 | -104.1464
20.8741 | -104.1319
20.8778 | -104.1193
20.8815 | -104.108
20.8852 | -104.0973
20.8889 | -104.0867
20.8926 | -104.0782
20.8963 | -104.0683
20.9 | -104.0579
20.9037 | -104.0465
20.9074 | -104.0369
20.9111 | -104.0286
20.9148 | -104.0219
20.9185 | -104.0165
20.9222 | -104.0096
20.9259 | -104.0025
20.9296 | -103.9944
20.9333 | -103.9884
20.937 | -103.9827
20.9407 | -103.9786
20.9444 | -103.9757
20.9481 | -103.9725
20.9519 | -103.9686
20.9556 | -103.966
20.9593 | -103.9627
20.963 | -103.9603
20.9667 | -103.9574
20.9704 | -103.9553
20.9741 | -103.9554
20.9778 | -103.9535
20.9815 | -103.9532
20.9852 | -103.9537
20.9889 | -103.9542
20.9926 | -103.9535
20.9963 | -103.9542
21 | -103.9585
21.0037 | -103.9619
21.0074 | -103.9642
21.0111 | -103.9685
21.0148 | -103.9732
21.0185 | -103.9792
21.0222 | -103.9843
21.0259 | -103.9905
21.0296 | -103.997
21.0333 | -104.0024
21.037 | -104.0082
21.0407 | -104.0186
21.0444 | -104.0318
21.0481 | -104.0435
21.0519 | -104.0505
21.0556 | -104.0565
21.0593 | -104.0643
21.063 | -104.0755
21.0667 | -104.0837
21.0704 | -104.0912
21.0741 | -104.1007
21.0778 | -104.1104
21.0815 | -104.119
21.0852 | -104.1231
21.0889 | -104.1253
21.0926 | -104.1289
21.0963 | -104.1311
21.1 | -104.1266
21.1037 | -104.1177
21.1074 | -104.1092
21.1111 | -104.1038
21.1148 | -104.0968
21.1185 | -104.0854
21.1222 | -104.0697
21.1259 | -104.0521
21.1296 | -104.0324
21.1333 | -104.0121
21.137 | -103.99
21.1407 | -103.9691
21.1444 | -103.9463
21.1481 | -103.9211
21.1519 | -103.8952
21.1556 | -103.8684
21.1593 | -103.837
21.163 | -103.8019
21.1667 | -103.7658
21.1704 | -103.729
21.1741 | -103.6923
21.1778 | -103.6537
21.1815 | -103.6133
21.1852 | -103.5733
21.1889 | -103.533
21.1926 | -103.49
21.1963 | -103.4493
21.2 | -103.4109
21.2037 | -103.3749
21.2074 | -103.3391
21.2111 | -103.3027
21.2148 | -103.2669
21.2185 | -103.2325
21.2222 | -103.2011
21.2259 | -103.1729
21.2296 | -103.1475
21.2333 | -103.1227
21.237 | -103.0987
21.2407 | -103.0729
21.2444 | -103.0467
21.2481 | -103.0251
21.2519 | -103.01
21.2556 | -102.9966
21.2593 | -102.9859
21.263 | -102.9767
21.2667 | -102.9676
21.2704 | -102.9597
21.2741 | -102.9541
21.2778 | -102.9506
21.2815 | -102.9457
21.2852 | -102.942
21.2889 | -102.9413
21.2926 | -102.9411
21.2963 | -102.9418
21.3 | -102.9419
21.3037 | -102.9416
21.3074 | -102.9419
21.3111 | -102.9399
21.3148 | -102.9382
21.3185 | -102.9381
21.3222 | -102.9391
21.3259 | -102.9389
21.3296 | -102.9385
21.3333 | -102.9396
21.337 | -102.9405
21.3407 | -102.9398
21.3444 | -102.9397
21.3481 | -102.9421
21.3519 | -102.947
21.3556 | -102.9508
21.3593 | -102.953
21.363 | -102.9544
21.3667 | -102.9549
21.3704 | -102.9628
21.3741 | -102.967
21.3778 | -102.9675
21.3815 | -102.9687
21.3852 | -102.975
21.3889 | -102.9859
21.3926 | -102.9974
21.3963 | -103.007
21.4 | -103.0172
21.4037 | -103.0327
21.4074 | -103.0481
21.4111 | -103.0566
21.4148 | -103.0625
21.4185 | -103.0707
21.4222 | -103.0825
21.4259 | -103.0967
21.4296 | -103.1101
21.4333 | -103.1235
21.437 | -103.136
21.4407 | -103.1492
21.4444 | -103.1606
21.4481 | -103.1703
21.4519 | -103.182
21.4556 | -103.1956
21.4593 | -103.21
21.463 | -103.2255
21.4667 | -103.2407
21.4704 | -103.2555
21.4741 | -103.2715
21.4778 | -103.2874
21.4815 | -103.3027
21.4852 | -103.3189
21.4889 | -103.3363
21.4926 | -103.3543
21.4963 | -103.3717
21.5 | -103.387
21.5037 | -103.4009
21.5074 | -103.4146
21.5111 | -103.431
21.5148 | -103.448
21.5185 | -103.4647
21.5222 | -103.4795
21.5259 | -103.4938
21.5296 | -103.5069
21.5333 | -103.5206
21.537 | -103.5363
21.5407 | -103.5511
21.5444 | -103.5687
21.5481 | -103.5876
21.5519 | -103.6034
21.5556 | -103.6125
21.5593 | -103.6201
21.563 | -103.6299
21.5667 | -103.6395
21.5704 | -103.6469
21.5741 | -103.6544
21.5778 | -103.6616
21.5815 | -103.6681
21.5852 | -103.6752
21.5889 | -103.6813
21.5926 | -103.6849
21.5963 | -103.6895
21.6 | -103.6934
21.6037 | -103.6962
21.6074 | -103.7014
21.6111 | -103.7076
21.6148 | -103.7146
21.6185 | -103.7197
21.6222 | -103.7234
21.6259 | -103.7275
21.6296 | -103.7331
21.6333 | -103.7387
21.637 | -103.7431
21.6407 | -103.7492
21.6444 | -103.7554
21.6481 | -103.7601
21.6519 | -103.7648
21.6556 | -103.7708
21.6593 | -103.778
21.663 | -103.7861
21.6667 | -103.7959
21.6704 | -103.8097
21.6741 | -103.82
21.6778 | -103.8286
21.6815 | -103.8409
21.6852 | -103.8569
21.6889 | -103.8719
21.6926 | -103.8842
21.6963 | -103.8962
21.7 | -103.9079
21.7037 | -103.9215
21.7074 | -103.935
21.7111 | -103.9459
21.7148 | -103.9595
21.7185 | -103.9772
21.7222 | -103.993
21.7259 | -104.0054
21.7296 | -104.017
21.7333 | -104.0273
21.737 | -104.0376
21.7407 | -104.05
21.7444 | -104.0635
21.7481 | -104.0731
21.7519 | -104.0821
21.7556 | -104.0937
21.7593 | -104.1037
21.763 | -104.1105
21.7667 | -104.1133
21.7704 | -104.1137
21.7741 | -104.1154
21.7778 | -104.1178
21.7815 | -104.1163
21.7852 | -104.1116
21.7889 | -104.1053
21.7926 | -104.0988
21.7963 | -104.0916
21.8 | -104.0845
21.8037 | -104.0783
21.8074 | -104.0702
21.8111 | -104.0614
21.8148 | -104.0484
21.8185 | -104.031
21.8222 | -104.0143
21.8259 | -103.9973
21.8296 | -103.9827
21.8333 | -103.969
21.837 | -103.9527
21.8407 | -103.9328
21.8444 | -103.9121
21.8481 | -103.8909
21.8519 | -103.8675
21.8556 | -103.8454
21.8593 | -103.8253
21.863 | -103.8068
21.8667 | -103.7891
21.8704 | -103.7693
21.8741 | -103.751
21.8778 | -103.7362
21.8815 | -103.7207
21.8852 | -103.7027
21.8889 | -103.6861
21.8926 | -103.6717
21.8963 | -103.6589
21.9 | -103.6479
21.9037 | -103.6371
21.9074 | -103.6367
21.9111 | -103.6282
21.9148 | -103.6224
21.9185 | -103.6156
21.9222 | -103.6082
21.9259 | -103.604
21.9296 | -103.6037
21.9333 | -103.6041
21.937 | -103.6043
21.9407 | -103.607
21.9444 | -103.6116
21.9481 | -103.6171
21.9519 | -103.6222
21.9556 | -103.6295
21.9593 | -103.6384
21.963 | -103.6451
21.9667 | -103.6537
21.9704 | -103.664
21.9741 | -103.6742
21.9778 | -103.6846
21.9815 | -103.6965
21.9852 | -103.7113
21.9889 | -103.7272
21.9926 | -103.7429
21.9963 | -103.7566
22 | -103.7708
22.0037 | -103.7863
22.0074 | -103.8031
22.0111 | -103.821
22.0148 | -103.838
22.0185 | -103.8564
22.0222 | -103.8764
22.0259 | -103.8939
22.0296 | -103.9101
22.0333 | -103.929
22.037 | -103.95
22.0407 | -103.9684
22.0444 | -103.9874
22.0481 | -104.0073
22.0519 | -104.0243
22.0556 | -104.0399
22.0593 | -104.0563
22.063 | -104.0723
22.0667 | -104.0875
22.0704 | -104.1015
22.0741 | -104.116
22.0778 | -104.1332
22.0815 | -104.1541
22.0852 | -104.1727
22.0889 | -104.1879
22.0926 | -104.2042
22.0963 | -104.2207
22.1 | -104.2377
22.1037 | -104.2522
22.1074 | -104.2637
22.1111 | -104.2739
22.1148 | -104.286
22.1185 | -104.2997
22.1222 | -104.3105
22.1259 | -104.3208
22.1296 | -104.3297
22.1333 | -104.3365
22.137 | -104.3444
22.1407 | -104.3513
22.1444 | -104.3551
22.1481 | -104.359
22.1519 | -104.3631
22.1556 | -104.3656
22.1593 | -104.3662
22.163 | -104.3649
22.1667 | -104.3639
22.1704 | -104.3618
22.1741 | -104.3578
22.1778 | -104.3513
22.1815 | -104.341
22.1852 | -104.3265
22.1889 | -104.3129
22.1926 | -104.3019
22.1963 | -104.2915
22.2 | -104.277
22.2037 | -104.2599
22.2074 | -104.2433
22.2111 | -104.2258
22.2148 | -104.2059
22.2185 | -104.1858
22.2222 | -104.1638
22.2259 | -104.1415
22.2296 | -104.1177
22.2333 | -104.0952
22.237 | -104.0714
22.2407 | -104.0462
22.2444 | -104.018
22.2481 | -103.9879
22.2519 | -103.9587
22.2556 | -103.929
22.2593 | -103.8988
22.263 | -103.8709
22.2667 | -103.8454
22.2704 | -103.8181
22.2741 | -103.7915
22.2778 | -103.7669
22.2815 | -103.7408
22.2852 | -103.7139
22.2889 | -103.6894
22.2926 | -103.6658
22.2963 | -103.6419
22.3 | -103.6185
22.3037 | -103.5952
22.3074 | -103.5742
22.3111 | -103.5572
22.3148 | -103.5416
22.3185 | -103.5259
22.3222 | -103.5111
22.3259 | -103.499
22.3296 | -103.4887
22.3333 | -103.4796
22.337 | -103.473
22.3407 | -103.4683
22.3444 | -103.4621
22.3481 | -103.4582
22.3519 | -103.4589
22.3556 | -103.4624
22.3593 | -103.4668
22.363 | -103.4698
22.3667 | -103.4734
22.3704 | -103.48
22.3741 | -103.4897
22.3778 | -103.5018
22.3815 | -103.5145
22.3852 | -103.528
22.3889 | -103.5431
22.3926 | -103.5582
22.3963 | -103.57
22.4 | -103.5795
22.4037 | -103.5892
22.4074 | -103.6006
22.4111 | -103.6159
22.4148 | -103.6294
22.4185 | -103.6414
22.4222 | -103.6551
22.4259 | -103.6681
22.4296 | -103.6787
22.4333 | -103.6871
22.437 | -103.6919
22.4407 | -103.6797
22.4444 | -103.6827
22.4481 | -103.6834
22.4519 | -103.6824
22.4556 | -103.6791
22.4593 | -103.6733
22.463 | -103.6672
22.4667 | -103.661
22.4704 | -103.6534
22.4741 | -103.645
22.4778 | -103.6341
22.4815 | -103.6236
22.4852 | -103.6101
22.4889 | -103.5946
22.4926 | -103.5802
22.4963 | -103.5656
22.5 | -103.551
22.5037 | -103.5367
22.5074 | -103.5191
22.5111 | -103.4995
22.5148 | -103.4805
22.5185 | -103.4658
22.5222 | -103.4514
22.5259 | -103.4351
22.5296 | -103.4144
22.5333 | -103.3945
22.537 | -103.3776
22.5407 | -103.364
22.5444 | -103.3497
22.5481 | -103.336
22.5519 | -103.3209
22.5556 | -103.3087
22.5593 | -103.2981
22.563 | -103.2877
22.5667 | -103.2769
22.5704 | -103.2676
22.5741 | -103.2599
22.5778 | -103.2523
22.5815 | -103.2409
22.5852 | -103.2285
22.5889 | -103.2177
22.5926 | -103.2089
22.5963 | -103.2022
22.6 | -103.1931
22.6037 | -103.1835
22.6074 | -103.1756
22.6111 | -103.1675
22.6148 | -103.1607
22.6185 | -103.1532
22.6222 | -103.1477
22.6259 | -103.1417
22.6296 | -103.1355
22.6333 | -103.1297
22.637 | -103.1226
22.6407 | -103.1167
22.6444 | -103.1113
22.6481 | -103.1062
22.6519 | -103.1026
22.6556 | -103.0991
22.6593 | -103.0979
22.663 | -103.0998
22.6667 | -103.1006
22.6704 | -103.0995
22.6741 | -103.0993
22.6778 | -103.1023
22.6815 | -103.1088
22.6852 | -103.115
22.6889 | -103.1211
22.6926 | -103.128
22.6963 | -103.1364
22.7 | -103.1461
22.7037 | -103.158
22.7074 | -103.17
22.7111 | -103.1835
22.7148 | -103.2008
22.7185 | -103.2223
22.7222 | -103.2444
22.7259 | -103.2644
22.7296 | -103.2836
22.7333 | -103.3077
22.737 | -103.3314
22.7407 | -103.3537
22.7444 | -103.3775
22.7481 | -103.4036
22.7519 | -103.4334
22.7556 | -103.4671
22.7593 | -103.5026
22.763 | -103.535
22.7667 | -103.5659
22.7704 | -103.6004
22.7741 | -103.6356
22.7778 | -103.6701
22.7815 | -103.7053
22.7852 | -103.7397
22.7889 | -103.7742
22.7926 | -103.8074
22.7963 | -103.8404
22.8 | -103.8738
22.8037 | -103.9063
22.8074 | -103.9385
22.8111 | -103.9697
22.8148 | -103.999
22.8185 | -104.0286
22.8222 | -104.0585
22.8259 | -104.0865
22.8296 | -104.1134
22.8333 | -104.1398
22.837 | -104.1643
22.8407 | -104.1843
22.8444 | -104.2025
22.8481 | -104.2172
22.8519 | -104.2329
22.8556 | -104.2481
22.8593 | -104.2619
22.863 | -104.2721
22.8667 | -104.2817
22.8704 | -104.2892
22.8741 | -104.298
22.8778 | -104.3015
22.8815 | -104.3021
22.8852 | -104.3035
22.8889 | -104.3045
22.8926 | -104.3011
22.8963 | -104.2962
22.9 | -104.2926
22.9037 | -104.2889
22.9074 | -104.2838
22.9111 | -104.2794
22.9148 | -104.2731
22.9185 | -104.2619
22.9222 | -104.2483
22.9259 | -104.2348
22.9296 | -104.2171
22.9333 | -104.1973
22.937 | -104.1781
22.9407 | -104.1616
22.9444 | -104.1448
22.9481 | -104.1265
22.9519 | -104.107
22.9556 | -104.0873
22.9593 | -104.0689
22.963 | -104.047
22.9667 | -104.0223
22.9704 | -103.9984
22.9741 | -103.9731
22.9778 | -103.9505
22.9815 | -103.9278
22.9852 | -103.9041
22.9889 | -103.8786
22.9926 | -103.8523
22.9963 | -103.8283
23 | -103.8044
23.0037 | -103.7802
23.0074 | -103.7574
23.0111 | -103.7381
23.0148 | -103.717
23.0185 | -103.6917
23.0222 | -103.6662
23.0259 | -103.6413
23.0296 | -103.6194
23.0333 | -103.5991
23.037 | -103.5804
23.0407 | -103.564
23.0444 | -103.5479
23.0481 | -103.5336
23.0519 | -103.5215
23.0556 | -103.5094
23.0593 | -103.4957
23.063 | -103.4855
23.0667 | -103.4769
23.0704 | -103.4706
23.0741 | -103.4651
23.0778 | -103.4591
23.0815 | -103.4553
23.0852 | -103.454
23.0889 | -103.454
23.0926 | -103.456
23.0963 | -103.4596
23.1 | -103.4616
23.1037 | -103.4643
23.1074 | -103.4701
23.1111 | -103.4792
23.1148 | -103.4882
23.1185 | -103.499
23.1222 | -103.5116
23.1259 | -103.5272
23.1296 | -103.5423
23.1333 | -103.5532
23.137 | -103.5627
23.1407 | -103.5773
23.1444 | -103.5935
23.1481 | -103.6091
23.1519 | -103.6265
23.1556 | -103.6443
23.1593 | -103.662
23.163 | -103.6806
23.1667 | -103.7
23.1704 | -103.7197
23.1741 | -103.7399
23.1778 | -103.7616
23.1815 | -103.7831
23.1852 | -103.8009
23.1889 | -103.818
23.1926 | -103.8356
23.1963 | -103.8528
23.2 | -103.8679
23.2037 | -103.8838
23.2074 | -103.8976
23.2111 | -103.9112
23.2148 | -103.9253
23.2185 | -103.9394
23.2222 | -103.9511
23.2259 | -103.9604
23.2296 | -103.9708
23.2333 | -103.9815
23.237 | -103.9882
23.2407 | -103.9937
23.2444 | -103.9988
23.2481 | -104.0044
23.2519 | -104.0095
23.2556 | -104.016
23.2593 | -104.0208
23.263 | -104.0238
23.2667 | -104.024
23.2704 | -104.0227
23.2741 | -104.0233
23.2778 | -104.0212
23.2815 | -104.018
23.2852 | -104.0155
23.2889 | -104.0147
23.2926 | -104.0154
23.2963 | -104.0159
23.3 | -104.0141
23.3037 | -104.0146
23.3074 | -104.014
23.3111 | -104.0135
23.3148 | -104.0128
23.3185 | -104.0149
23.3222 | -104.0166
23.3259 | -104.0155
23.3296 | -104.0146
23.3333 | -104.0147
23.337 | -104.0151
23.3407 | -104.0151
23.3444 | -104.0172
23.3481 | -104.0205
23.3519 | -104.023
23.3556 | -104.0251
23.3593 | -104.0243
23.363 | -104.0241
23.3667 | -104.0248
23.3704 | -104.0215
23.3741 | -104.0186
23.3778 | -104.0165
23.3815 | -104.0164
23.3852 | -104.0156
23.3889 | -104.0157
23.3926 | -104.0168
23.3963 | -104.0178
23.4 | -104.0161
23.4037 | -104.0146
23.4074 | -104.0134
23.4111 | -104.0128
23.4148 | -104.0118
23.4185 | -104.0119
23.4222 | -104.0116
23.4259 | -104.0104
23.4296 | -104.0134
23.4333 | -104.0164
23.437 | -104.0153
23.4407 | -104.0146
23.4444 | -104.0168
23.4481 | -104.0193
23.4519 | -104.0209
23.4556 | -104.0236
23.4593 | -104.0286
23.463 | -104.0364
23.4667 | -104.0464
23.4704 | -104.0535
23.4741 | -104.0591
23.4778 | -104.0668
23.4815 | -104.073
23.4852 | -104.0776
23.4889 | -104.0829
23.4926 | -104.0897
23.4963 | -104.0973
23.5 | -104.1018
23.5037 | -104.1087
23.5074 | -104.1199
23.5111 | -104.1271
23.5148 | -104.1345
23.5185 | -104.1412
23.5222 | -104.1468
23.5259 | -104.1534
23.5296 | -104.1602
23.5333 | -104.1611
23.537 | -104.1596
23.5407 | -104.164
23.5444 | -104.1699
23.5481 | -104.1708
23.5519 | -104.1688
23.5556 | -104.1684
23.5593 | -104.17
23.563 | -104.1709
23.5667 | -104.1686
23.5704 | -104.1663
23.5741 | -104.1619
23.5778 | -104.1587
23.5815 | -104.1583
23.5852 | -104.1522
23.5889 | -104.1361
23.5926 | -104.1206
23.5963 | -104.1061
23.6 | -104.0882
23.6037 | -104.0629
23.6074 | -104.0373
23.6111 | -104.0187
23.6148 | -104.0013
23.6185 | -103.981
23.6222 | -103.9601
23.6259 | -103.9358
23.6296 | -103.9075
23.6333 | -103.8794
23.637 | -103.851
23.6407 | -103.8217
23.6444 | -103.7909
23.6481 | -103.7623
23.6519 | -103.7324
23.6556 | -103.6987
23.6593 | -103.6654
23.663 | -103.6358
23.6667 | -103.6042
23.6704 | -103.572
23.6741 | -103.5427
23.6778 | -103.5132
23.6815 | -103.4849
23.6852 | -103.4576
23.6889 | -103.4312
23.6926 | -103.403
23.6963 | -103.3729
23.7 | -103.3449
23.7037 | -103.3206
23.7074 | -103.2947
23.7111 | -103.2674
23.7148 | -103.2436
23.7185 | -103.2247
23.7222 | -103.2073
23.7259 | -103.1898
23.7296 | -103.1749
23.7333 | -103.1617
23.737 | -103.1495
23.7407 | -103.139
23.7444 | -103.1249
23.7481 | -103.1097
23.7519 | -103.0968
23.7556 | -103.0877
23.7593 | -103.0828
23.763 | -103.0795
23.7667 | -103.0769
23.7704 | -103.0782
23.7741 | -103.0807
23.7778 | -103.0821
23.7815 | -103.0828
23.7852 | -103.0819
23.7889 | -103.0812
23.7926 | -103.0841
23.7963 | -103.0885
23.8 | -103.092
23.8037 | -103.0941
23.8074 | -103.0964
23.8111 | -103.0987
23.8148 | -103.1027
23.8185 | -103.1084
23.8222 | -103.1143
23.8259 | -103.1209
23.8296 | -103.1276
23.8333 | -103.1357
23.837 | -103.1406
23.8407 | -103.1444
23.8444 | -103.1471
23.8481 | -103.1513
23.8519 | -103.156
23.8556 | -103.1605
23.8593 | -103.1645
23.863 | -103.1673
23.8667 | -103.1699
23.8704 | -103.1759
23.8741 | -103.1826
23.8778 | -103.1856
23.8815 | -103.1838
23.8852 | -103.1852
23.8889 | -103.1896
23.8926 | -103.1909
23.8963 | -103.1889
23.9 | -103.1876
23.9037 | -103.1881
23.9074 | -103.1896
23.9111 | -103.1924
23.9148 | -103.1941
23.9185 | -103.1926
23.9222 | -103.1918
23.9259 | -103.1984
23.9296 | -103.2053
23.9333 | -103.2077
23.937 | -103.2105
23.9407 | -103.218
23.9444 | -103.2255
23.9481 | -103.2292
23.9519 | -103.2306
23.9556 | -103.2326
23.9593 | -103.2351
23.963 | -103.242
23.9667 | -103.2489
23.9704 | -103.2524
23.9741 | -103.256
23.9778 | -103.2643
23.9815 | -103.2754
23.9852 | -103.2837
23.9889 | -103.2921
23.9926 | -103.3046
23.9963 | -103.3168
24 | -103.326
24.0037 | -103.3366
24.0074 | -103.3472
24.0111 | -103.3542
24.0148 | -103.3641
24.0185 | -103.374
24.0222 | -103.3816
24.0259 | -103.3891
24.0296 | -103.3999
24.0333 | -103.413
24.037 | -103.4241
24.0407 | -103.4327
24.0444 | -103.439
24.0481 | -103.4503
24.0519 | -103.463
24.0556 | -103.4701
24.0593 | -103.4748
24.063 | -103.4799
24.0667 | -103.4865
24.0704 | -103.4936
24.0741 | -103.5002
24.0778 | -103.5073
24.0815 | -103.5138
24.0852 | -103.525
24.0889 | -103.5381
24.0926 | -103.5493
24.0963 | -103.5581
24.1 | -103.5631
24.1037 | -103.568
24.1074 | -103.5748
24.1111 | -103.5811
24.1148 | -103.5861
24.1185 | -103.5935
24.1222 | -103.604
24.1259 | -103.6188
24.1296 | -103.6334
24.1333 | -103.6458
24.137 | -103.6597
24.1407 | -103.6743
24.1444 | -103.6877
24.1481 | -103.7031
24.1519 | -103.7196
24.1556 | -103.7351
24.1593 | -103.751
24.163 | -103.7668
24.1667 | -103.7818
24.1704 | -103.7979
24.1741 | -103.8168
24.1778 | -103.8368
24.1815 | -103.8552
24.1852 | -103.8744
24.1889 | -103.8936
24.1926 | -103.9115
24.1963 | -103.9315
24.2 | -103.9528
24.2037 | -103.9741
24.2074 | -103.9959
24.2111 | -104.0186
24.2148 | -104.0406
24.2185 | -104.0608
24.2222 | -104.0832
24.2259 | -104.1054
24.2296 | -104.1229
24.2333 | -104.1389
24.237 | -104.1571
24.2407 | -104.1761
24.2444 | -104.1928
24.2481 | -104.203
24.2519 | -104.2112
24.2556 | -104.223
24.2593 | -104.2348
24.263 | -104.2451
24.2667 | -104.2497
24.2704 | -104.2519
24.2741 | -104.256
24.2778 | -104.2586
24.2815 | -104.2566
24.2852 | -104.2465
24.2889 | -104.2311
24.2926 | -104.2121
24.2963 | -104.1932
24.3 | -104.169
24.3037 | -104.147
24.3074 | -104.1269
24.3111 | -104.1102
24.3148 | -104.0857
24.3185 | -104.0544
24.3222 | -104.021
24.3259 | -103.9874
24.3296 | -103.9527
24.3333 | -103.9163
24.337 | -103.8787
24.3407 | -103.8393
24.3444 | -103.7973
24.3481 | -103.7533
24.3519 | -103.7093
24.3556 | -103.6607
24.3593 | -103.6094
24.363 | -103.5575
24.3667 | -103.5052
24.3704 | -103.4533
24.3741 | -103.4015
24.3778 | -103.3486
24.3815 | -103.2981
24.3852 | -103.2468
24.3889 | -103.1951
24.3926 | -103.1427
24.3963 | -103.0902
24.4 | -103.0401
24.4037 | -102.9931
24.4074 | -102.9488
24.4111 | -102.9049
24.4148 | -102.8614
24.4185 | -102.8199
24.4222 | -102.7808
24.4259 | -102.7443
24.4296 | -102.7107
24.4333 | -102.6811
24.437 | -102.6545
24.4407 | -102.6289
24.4444 | -102.6067
24.4481 | -102.5876
24.4519 | -102.5692
24.4556 | -102.5506
24.4593 | -102.5369
24.463 | -102.5282
24.4667 | -102.5221
24.4704 | -102.5183
24.4741 | -102.5175
24.4778 | -102.5186
24.4815 | -102.5217
24.4852 | -102.5274
24.4889 | -102.5358
24.4926 | -102.5465
24.4963 | -102.5578
24.5 | -102.5718
24.5037 | -102.5876
24.5074 | -102.6042
24.5111 | -102.6197
24.5148 | -102.6373
24.5185 | -102.6576
24.5222 | -102.6788
24.5259 | -102.6999
24.5296 | -102.7228
24.5333 | -102.7437
24.537 | -102.7657
24.5407 | -102.789
24.5444 | -102.8099
24.5481 | -102.826
24.5519 | -102.8403
24.5556 | -102.8519
24.5593 | -102.8629
24.563 | -102.8717
24.5667 | -102.8751
24.5704 | -102.8788
24.5741 | -102.8827
24.5778 | -102.8199
24.5815 | -102.7913
24.5852 | -102.7602
24.5889 | -102.7249
24.5926 | -102.6899
24.5963 | -102.6558
24.6 | -102.6249
24.6037 | -102.5881
24.6074 | -102.544
24.6111 | -102.5
24.6148 | -102.455
24.6185 | -102.405
24.6222 | -102.3506
24.6259 | -102.2945
24.6296 | -102.237
24.6333 | -102.1775
24.637 | -102.1164
24.6407 | -102.0547
24.6444 | -101.9876
24.6481 | -101.9175
24.6519 | -101.8403
24.6556 | -101.762
24.6593 | -101.6818
24.663 | -101.6033
24.6667 | -101.5251
24.6704 | -101.4455
24.6741 | -101.3663
24.6778 | -101.2888
24.6815 | -101.2132
24.6852 | -101.1383
24.6889 | -101.0649
24.6926 | -100.995
24.6963 | -100.9285
24.7 | -100.8645
24.7037 | -100.804
24.7074 | -100.7488
24.7111 | -100.695
24.7148 | -100.6423
24.7185 | -100.5943
24.7222 | -100.5508
24.7259 | -100.512
24.7296 | -100.4785
24.7333 | -100.4502
24.737 | -100.4271
24.7407 | -100.4085
24.7444 | -100.3941
24.7481 | -100.3853
24.7519 | -100.3802
24.7556 | -100.3802
24.7593 | -100.3872
24.763 | -100.3992
24.7667 | -100.4161
24.7704 | -100.4382
24.7741 | -100.4654
24.7778 | -100.4971
24.7815 | -100.5331
24.7852 | -100.5731
24.7889 | -100.6179
24.7926 | -100.668
24.7963 | -100.7225
24.8 | -100.7807
24.8037 | -100.8446
24.8074 | -100.911
24.8111 | -100.9767
24.8148 | -101.0466
24.8185 | -101.1204
24.8222 | -101.197
24.8259 | -101.2761
24.8296 | -101.3579
24.8333 | -101.4404
24.837 | -101.5225
24.8407 | -101.6033
24.8444 | -101.6853
24.8481 | -101.7675
24.8519 | -101.8502
24.8556 | -101.9301
24.8593 | -102.0115
24.863 | -102.0904
24.8667 | -102.165
24.8704 | -102.2399
24.8741 | -102.3147
24.8778 | -102.3851
24.8815 | -102.4529
24.8852 | -102.5174
24.8889 | -102.5784
24.8926 | -102.6374
24.8963 | -102.6943
24.9 | -102.7474
24.9037 | -102.7998
24.9074 | -102.8469
24.9111 | -102.8903
24.9148 | -102.9342
24.9185 | -102.9725
24.9222 | -103.0065
24.9259 | -103.0464
24.9296 | -103.099
24.9333 | -103.1381
24.937 | -103.1664
24.9407 | -103.1888
24.9444 | -103.2034
24.9481 | -103.2113
24.9519 | -103.2105
24.9556 | -103.2069
24.9593 | -103.2071
24.963 | -103.2033
24.9667 | -103.1955
24.9704 | -103.1882
24.9741 | -103.1814
24.9778 | -103.1742
24.9815 | -103.1666
24.9852 | -103.1595
24.9889 | -103.1517
24.9926 | -103.1426
24.9963 | -103.1362
25 | -103.133
25.0037 | -103.1297
25.0074 | -103.1267
25.0111 | -103.125
25.0148 | -103.1236
25.0185 | -103.1219
25.0222 | -103.1233
25.0259 | -103.1249
25.0296 | -103.1244
25.0333 | -103.1252
25.037 | -103.1283
25.0407 | -103.133
25.0444 | -103.1384
25.0481 | -103.143
25.0519 | -103.1474
25.0556 | -103.1544
25.0593 | -103.1623
25.063 | -103.1754
25.0667 | -103.191
25.0704 | -103.201
25.0741 | -103.21
25.0778 | -103.2221
25.0815 | -103.2322
25.0852 | -103.2428
25.0889 | -103.2548
25.0926 | -103.2674
25.0963 | -103.2831
25.1 | -103.297
25.1037 | -103.3113
25.1074 | -103.3255
25.1111 | -103.3429
25.1148 | -103.3564
25.1185 | -103.3696
25.1222 | -103.3839
25.1259 | -103.3975
25.1296 | -103.4111
25.1333 | -103.4256
25.137 | -103.4401
25.1407 | -103.4555
25.1444 | -103.4726
25.1481 | -103.4871
25.1519 | -103.5009
25.1556 | -103.5148
25.1593 | -103.5283
25.163 | -103.5425
25.1667 | -103.5598
25.1704 | -103.5756
25.1741 | -103.5891
25.1778 | -103.6012
25.1815 | -103.6135
25.1852 | -103.6249
25.1889 | -103.6392
25.1926 | -103.6566
25.1963 | -103.6742
25.2 | -103.6914
25.2037 | -103.7069
25.2074 | -103.7189
25.2111 | -103.7341
25.2148 | -103.747
25.2185 | -103.7613
25.2222 | -103.7754
25.2259 | -103.7881
25.2296 | -103.8006
25.2333 | -103.8126
25.237 | -103.8271
25.2407 | -103.8408
25.2444 | -103.8526
25.2481 | -103.8633
25.2519 | -103.8762
25.2556 | -103.8898
25.2593 | -103.9011
25.263 | -103.9125
25.2667 | -103.9226
25.2704 | -103.9298
25.2741 | -103.9372
25.2778 | -103.9444
25.2815 | -103.9508
25.2852 | -103.957
25.2889 | -103.9647
25.2926 | -103.9739
25.2963 | -103.9796
25.3 | -103.9831
25.3037 | -103.9846
25.3074 | -103.9834
25.3111 | -103.9847
25.3148 | -103.9851
25.3185 | -103.9815
25.3222 | -103.9767
25.3259 | -103.9738
25.3296 | -103.9685
25.3333 | -103.9627
25.337 | -103.9571
25.3407 | -103.9513
25.3444 | -103.9445
25.3481 | -103.9378
25.3519 | -103.9305
25.3556 | -103.921
25.3593 | -103.9117
25.363 | -103.9003
25.3667 | -103.8892
25.3704 | -103.8795
25.3741 | -103.8693
25.3778 | -103.8577
25.3815 | -103.8446
25.3852 | -103.8318
25.3889 | -103.8174
25.3926 | -103.8023
25.3963 | -103.7883
25.4 | -103.7745
25.4037 | -103.7621
25.4074 | -103.7511
25.4111 | -103.7362
25.4148 | -103.7194
25.4185 | -103.7033
25.4222 | -103.6892
25.4259 | -103.6768
25.4296 | -103.6643
25.4333 | -103.6497
25.437 | -103.6351
25.4407 | -103.6206
25.4444 | -103.6072
25.4481 | -103.5962
25.4519 | -103.5833
25.4556 | -103.5709
25.4593 | -103.5589
25.463 | -103.5471
25.4667 | -103.5346
25.4704 | -103.5236
25.4741 | -103.5125
25.4778 | -103.5007
25.4815 | -103.4909
25.4852 | -103.4834
25.4889 | -103.4745
25.4926 | -103.4632
25.4963 | -103.4557
25.5 | -103.4494
25.5037 | -103.4414
25.5074 | -103.4333
25.5111 | -103.4253
25.5148 | -103.4199
25.5185 | -103.4164
25.5222 | -103.412
25.5259 | -103.4064
25.5296 | -103.3988
25.5333 | -103.3926
25.537 | -103.39
25.5407 | -103.3872
25.5444 | -103.3834
25.5481 | -103.3777
25.5519 | -103.3737
25.5556 | -103.3717
25.5593 | -103.3694
25.563 | -103.3676
25.5667 | -103.3652
25.5704 | -103.364
25.5741 | -103.3646
25.5778 | -103.3638
25.5815 | -103.365
25.5852 | -103.3658
25.5889 | -103.3643
25.5926 | -103.3634
25.5963 | -103.3632
25.6 | -103.3622
25.6037 | -103.3615
25.6074 | -103.3636
25.6111 | -103.3661
25.6148 | -103.3689
25.6185 | -103.3716
25.6222 | -103.3744
25.6259 | -103.3785
25.6296 | -103.3829
25.6333 | -103.385
25.637 | -103.3861
25.6407 | -103.3878
25.6444 | -103.3895
25.6481 | -103.393
25.6519 | -103.3971
25.6556 | -103.4005
25.6593 | -103.4047
25.663 | -103.4097
25.6667 | -103.4169
25.6704 | -103.4223
25.6741 | -103.4251
25.6778 | -103.4281
25.6815 | -103.4329
25.6852 | -103.4382
25.6889 | -103.4421
25.6926 | -103.4472
25.6963 | -103.4516
25.7 | -103.4566
25.7037 | -103.4619
25.7074 | -103.4679
25.7111 | -103.4741
25.7148 | -103.4802
25.7185 | -103.4865
25.7222 | -103.4944
25.7259 | -103.503
25.7296 | -103.5108
25.7333 | -103.5177
25.737 | -103.5237
25.7407 | -103.5282
25.7444 | -103.5363
25.7481 | -103.5453
25.7519 | -103.553
25.7556 | -103.5593
25.7593 | -103.5663
25.763 | -103.571
25.7667 | -103.5761
25.7704 | -103.583
25.7741 | -103.5923
25.7778 | -103.6024
25.7815 | -103.6091
25.7852 | -103.6145
25.7889 | -103.6212
25.7926 | -103.6271
25.7963 | -103.6341
25.8 | -103.6426
25.8037 | -103.6505
25.8074 | -103.6579
25.8111 | -103.6656
25.8148 | -103.6746
25.8185 | -103.6827
25.8222 | -103.6904
25.8259 | -103.6986
25.8296 | -103.7065
25.8333 | -103.7135
25.837 | -103.7227
25.8407 | -103.7347
25.8444 | -103.745
25.8481 | -103.7538
25.8519 | -103.7641
25.8556 | -103.7747
25.8593 | -103.7828
25.863 | -103.7922
25.8667 | -103.8034
25.8704 | -103.8153
25.8741 | -103.8272
25.8778 | -103.8386
25.8815 | -103.8476
25.8852 | -103.858
25.8889 | -103.8716
25.8926 | -103.8851
25.8963 | -103.8946
25.9 | -103.9022
25.9037 | -103.9118
25.9074 | -103.9244
25.9111 | -103.9394
25.9148 | -103.9522
25.9185 | -103.9631
25.9222 | -103.9723
25.9259 | -103.9821
25.9296 | -103.9952
25.9333 | -104.0069
25.937 | -104.0161
25.9407 | -104.0251
25.9444 | -104.0332
25.9481 | -104.0421
25.9519 | -104.0519
25.9556 | -104.0605
25.9593 | -104.0698
25.963 | -104.0804
25.9667 | -104.0924
25.9704 | -104.1032
25.9741 | -104.1116
25.9778 | -104.1192
25.9815 | -104.1259
25.9852 | -104.1332
25.9889 | -104.1396
25.9926 | -104.1455
25.9963 | -104.1521
26 | -104.1615
26.0037 | -104.1715
26.0074 | -104.1814
26.0111 | -104.1901
26.0148 | -104.1958
26.0185 | -104.1996
26.0222 | -104.203
26.0259 | -104.2081
26.0296 | -104.2136
26.0333 | -104.2171
26.037 | -104.2208
26.0407 | -104.2265
26.0444 | -104.2317
26.0481 | -104.2364
26.0519 | -104.2403
26.0556 | -104.2427
26.0593 | -104.2462
26.063 | -104.2507
26.0667 | -104.2547
26.0704 | -104.2586
26.0741 | -104.2638
26.0778 | -104.2695
26.0815 | -104.2737
26.0852 | -104.2787
26.0889 | -104.283
26.0926 | -104.288
26.0963 | -104.2943
26.1 | -104.3015
26.1037 | -104.3112
26.1074 | -104.3222
26.1111 | -104.3305
26.1148 | -104.3389
26.1185 | -104.3486
26.1222 | -104.3577
26.1259 | -104.3642
26.1296 | -104.3738
26.1333 | -104.3841
26.137 | -104.3936
26.1407 | -104.402
26.1444 | -104.4099
26.1481 | -104.4201
26.1519 | -104.4305
26.1556 | -104.4383
26.1593 | -104.4479
26.163 | -104.458
26.1667 | -104.4681
26.1704 | -104.4783
26.1741 | -104.4871
26.1778 | -104.495
26.1815 | -104.5001
26.1852 | -104.5067
26.1889 | -104.5139
26.1926 | -104.5214
26.1963 | -104.527
26.2 | -104.5315
26.2037 | -104.5363
26.2074 | -104.5391
26.2111 | -104.5402
26.2148 | -104.5414
26.2185 | -104.5422
26.2222 | -104.5413
26.2259 | -104.5403
26.2296 | -104.5388
26.2333 | -104.5357
26.237 | -104.532
26.2407 | -104.5292
26.2444 | -104.5267
26.2481 | -104.5219
26.2519 | -104.5156
26.2556 | -104.5068
26.2593 | -104.4975
26.263 | -104.49
26.2667 | -104.4849
26.2704 | -104.4769
26.2741 | -104.4667
26.2778 | -104.4534
26.2815 | -104.4378
26.2852 | -104.4225
26.2889 | -104.4085
26.2926 | -104.3968
26.2963 | -104.3844
26.3 | -104.3689
26.3037 | -104.3531
26.3074 | -104.3402
26.3111 | -104.3272
26.3148 | -104.313
26.3185 | -104.2977
26.3222 | -104.2843
26.3259 | -104.2726
26.3296 | -104.2582
26.3333 | -104.2445
26.337 | -104.2293
26.3407 | -104.2152
26.3444 | -104.2041
26.3481 | -104.1926
26.3519 | -104.1814
26.3556 | -104.1708
26.3593 | -104.1614
26.363 | -104.1524
26.3667 | -104.1416
26.3704 | -104.131
26.3741 | -104.1212
26.3778 | -104.1131
26.3815 | -104.1068
26.3852 | -104.1029
26.3889 | -104.0985
26.3926 | -104.0939
26.3963 | -104.0894
26.4 | -104.0869
26.4037 | -104.0867
26.4074 | -104.0876
26.4111 | -104.0885
26.4148 | -104.0891
26.4185 | -104.0893
26.4222 | -104.0904
26.4259 | -104.0923
26.4296 | -104.0942
26.4333 | -104.0949
26.437 | -104.0965
26.4407 | -104.1001
26.4444 | -104.1037
26.4481 | -104.1061
26.4519 | -104.1097
26.4556 | -104.1156
26.4593 | -104.1225
26.463 | -104.1292
26.4667 | -104.1351
26.4704 | -104.1418
26.4741 | -104.1495
26.4778 | -104.1585
26.4815 | -104.1684
26.4852 | -104.1766
26.4889 | -104.1841
26.4926 | -104.1909
26.4963 | -104.1995
26.5 | -104.2071
26.5037 | -104.2125
26.5074 | -104.2179
26.5111 | -104.2257
26.5148 | -104.2335
26.5185 | -104.2406
26.5222 | -104.2473
26.5259 | -104.2576
26.5296 | -104.2666
26.5333 | -104.2729
26.537 | -104.2762
26.5407 | -104.2792
26.5444 | -104.282
26.5481 | -104.2858
26.5519 | -104.2905
26.5556 | -104.2931
26.5593 | -104.2955
26.563 | -104.2989
26.5667 | -104.3017
26.5704 | -104.3022
26.5741 | -104.3028
26.5778 | -104.3025
26.5815 | -104.3033
26.5852 | -104.3057
26.5889 | -104.3065
26.5926 | -104.3066
26.5963 | -104.3082
26.6 | -104.3099
26.6037 | -104.3087
26.6074 | -104.3054
26.6111 | -104.3031
26.6148 | -104.2995
26.6185 | -104.2958
26.6222 | -104.2908
26.6259 | -104.2837
26.6296 | -104.2794
26.6333 | -104.2765
26.637 | -104.2695
26.6407 | -104.261
26.6444 | -104.2563
26.6481 | -104.2517
26.6519 | -104.2478
26.6556 | -104.2443
26.6593 | -104.2399
26.663 | -104.2346
26.6667 | -104.2283
26.6704 | -104.2211
26.6741 | -104.215
26.6778 | -104.2087
26.6815 | -104.2014
26.6852 | -104.194
26.6889 | -104.1866
26.6926 | -104.1809
26.6963 | -104.1748
26.7 | -104.1694
26.7037 | -104.1644
26.7074 | -104.1567
26.7111 | -104.1479
26.7148 | -104.1415
26.7185 | -104.1348
26.7222 | -104.1283
26.7259 | -104.1219
26.7296 | -104.1159
26.7333 | -104.1091
26.737 | -104.1031
26.7407 | -104.0993
26.7444 | -104.0967
26.7481 | -104.0915
26.7519 | -104.0877
26.7556 | -104.0823
26.7593 | -104.075
26.763 | -104.0702
26.7667 | -104.0677
26.7704 | -104.0636
26.7741 | -104.059
26.7778 | -104.0544
26.7815 | -104.0481
26.7852 | -104.0427
26.7889 | -104.039
26.7926 | -104.035
26.7963 | -104.0318
26.8 | -104.0308
26.8037 | -104.0297
26.8074 | -104.0266
26.8111 | -104.0236
26.8148 | -104.0202
26.8185 | -104.0175
26.8222 | -104.0145
26.8259 | -104.0131
26.8296 | -104.0117
26.8333 | -104.0063
26.837 | -104.0001
26.8407 | -103.9953
26.8444 | -103.9907
26.8481 | -103.9866
26.8519 | -103.9818
26.8556 | -103.9764
26.8593 | -103.9699
26.863 | -103.9646
26.8667 | -103.9606
26.8704 | -103.957
26.8741 | -103.9517
26.8778 | -103.9453
26.8815 | -103.9393
26.8852 | -103.9338
26.8889 | -103.9266
26.8926 | -103.9169
26.8963 | -103.9084
26.9 | -103.9016
26.9037 | -103.8922
26.9074 | -103.8823
26.9111 | -103.8739
26.9148 | -103.8674
26.9185 | -103.8607
26.9222 | -103.8541
26.9259 | -103.8468
26.9296 | -103.8374
26.9333 | -103.8285
26.937 | -103.8191
26.9407 | -103.8095
26.9444 | -103.8004
26.9481 | -103.7932
26.9519 | -103.7873
26.9556 | -103.7805
26.9593 | -103.7728
26.963 | -103.761
26.9667 | -103.7516
26.9704 | -103.7434
26.9741 | -103.7333
26.9778 | -103.7239
26.9815 | -103.7169
26.9852 | -103.7084
26.9889 | -103.6979
26.9926 | -103.6882
26.9963 | -103.6789
27 | -103.6724
27.0037 | -103.6682
27.0074 | -103.6618
27.0111 | -103.6523
27.0148 | -103.6445
27.0185 | -103.6368
27.0222 | -103.6272
27.0259 | -103.6199
27.0296 | -103.612
27.0333 | -103.6033
27.037 | -103.5939
27.0407 | -103.585
27.0444 | -103.5768
27.0481 | -103.5709
27.0519 | -103.5657
27.0556 | -103.5615
27.0593 | -103.5561
27.063 | -103.551
27.0667 | -103.5469
27.0704 | -103.542
27.0741 | -103.5355
27.0778 | -103.5288
27.0815 | -103.5236
27.0852 | -103.5212
27.0889 | -103.5199
27.0926 | -103.5182
27.0963 | -103.5153
27.1 | -103.5131
27.1037 | -103.5121
27.1074 | -103.5118
27.1111 | -103.5125
27.1148 | -103.5145
27.1185 | -103.5156
27.1222 | -103.5159
27.1259 | -103.5169
27.1296 | -103.5175
27.1333 | -103.5172
27.137 | -103.5205
27.1407 | -103.5252
27.1444 | -103.5272
27.1481 | -103.5288
27.1519 | -103.5335
27.1556 | -103.5374
27.1593 | -103.5411
27.163 | -103.5465
27.1667 | -103.5522
27.1704 | -103.5578
27.1741 | -103.565
27.1778 | -103.5726
27.1815 | -103.5781
27.1852 | -103.5836
27.1889 | -103.5903
27.1926 | -103.5994
27.1963 | -103.6063
27.2 | -103.6123
27.2037 | -103.6195
27.2074 | -103.6261
27.2111 | -103.6324
27.2148 | -103.6415
27.2185 | -103.6516
27.2222 | -103.6596
27.2259 | -103.6679
27.2296 | -103.6764
27.2333 | -103.6857
27.237 | -103.6944
27.2407 | -103.7045
27.2444 | -103.7145
27.2481 | -103.723
27.2519 | -103.7319
27.2556 | -103.7415
27.2593 | -103.7493
27.263 | -103.7581
27.2667 | -103.7671
27.2704 | -103.7747
27.2741 | -103.7837
27.2778 | -103.794
27.2815 | -103.8038
27.2852 | -103.8119
27.2889 | -103.8211
27.2926 | -103.8303
27.2963 | -103.8381
27.3 | -103.847
27.3037 | -103.8565
27.3074 | -103.865
27.3111 | -103.8726
27.3148 | -103.8821
27.3185 | -103.8918
27.3222 | -103.9019
27.3259 | -103.9112
27.3296 | -103.9188
27.3333 | -103.9261
27.337 | -103.9342
27.3407 | -103.9436
27.3444 | -103.953
27.3481 | -103.9649
27.3519 | -103.9744
27.3556 | -103.9829
27.3593 | -103.9935
27.363 | -104.0044
27.3667 | -104.0145
27.3704 | -104.0266
27.3741 | -104.038
27.3778 | -104.0498
27.3815 | -104.0618
27.3852 | -104.0734
27.3889 | -104.0846
27.3926 | -104.0953
27.3963 | -104.1074
27.4 | -104.1196
27.4037 | -104.1331
27.4074 | -104.1455
27.4111 | -104.1537
27.4148 | -104.1637
27.4185 | -104.1752
27.4222 | -104.1892
27.4259 | -104.202
27.4296 | -104.2126
27.4333 | -104.2216
27.437 | -104.2294
27.4407 | -104.2367
27.4444 | -104.2453
27.4481 | -104.254
27.4519 | -104.2625
27.4556 | -104.2688
27.4593 | -104.2781
27.463 | -104.2874
27.4667 | -104.2943
27.4704 | -104.3002
27.4741 | -104.3073
27.4778 | -104.3119
27.4815 | -104.3155
27.4852 | -104.3201
27.4889 | -104.3216
27.4926 | -104.3215
27.4963 | -104.3211
27.5 | -104.3209
27.5037 | -104.322
27.5074 | -104.3214
27.5111 | -104.3209
27.5148 | -104.3227
27.5185 | -104.3243
27.5222 | -104.3221
27.5259 | -104.3195
27.5296 | -104.318
27.5333 | -104.3155
27.537 | -104.3113
27.5407 | -104.3069
27.5444 | -104.3024
27.5481 | -104.2968
27.5519 | -104.2899
27.5556 | -104.2794
27.5593 | -104.2683
27.563 | -104.2584
27.5667 | -104.2473
27.5704 | -104.2351
27.5741 | -104.2253
27.5778 | -104.2148
27.5815 | -104.2054
27.5852 | -104.1952
27.5889 | -104.183
27.5926 | -104.1704
27.5963 | -104.1585
27.6 | -104.1447
27.6037 | -104.1301
27.6074 | -104.118
27.6111 | -104.105
27.6148 | -104.0885
27.6185 | -104.0737
27.6222 | -104.0604
27.6259 | -104.0461
27.6296 | -104.0306
27.6333 | -104.0155
27.637 | -104.004
27.6407 | -103.9913
27.6444 | -103.9763
27.6481 | -103.9617
27.6519 | -103.9482
27.6556 | -103.9349
27.6593 | -103.9221
27.663 | -103.9083
27.6667 | -103.8963
27.6704 | -103.8835
27.6741 | -103.8714
27.6778 | -103.8612
27.6815 | -103.8519
27.6852 | -103.8429
27.6889 | -103.8328
27.6926 | -103.8224
27.6963 | -103.8122
27.7 | -103.8042
27.7037 | -103.7944
27.7074 | -103.7864
27.7111 | -103.7791
27.7148 | -103.7727
27.7185 | -103.7679
27.7222 | -103.7638
27.7259 | -103.7563
27.7296 | -103.7486
27.7333 | -103.7416
27.737 | -103.7361
27.7407 | -103.7316
27.7444 | -103.7288
27.7481 | -103.7245
27.7519 | -103.7189
27.7556 | -103.7148
27.7593 | -103.712
27.763 | -103.7112
27.7667 | -103.7089
27.7704 | -103.7063
27.7741 | -103.7031
27.7778 | -103.6991
27.7815 | -103.7
27.7852 | -103.6997
27.7889 | -103.6983
27.7926 | -103.6988
27.7963 | -103.6985
27.8 | -103.6969
27.8037 | -103.698
27.8074 | -103.6995
27.8111 | -103.7009
27.8148 | -103.7007
27.8185 | -103.7019
27.8222 | -103.7045
27.8259 | -103.7055
27.8296 | -103.7071
27.8333 | -103.7095
27.837 | -103.7121
27.8407 | -103.7142
27.8444 | -103.715
27.8481 | -103.7174
27.8519 | -103.72
27.8556 | -103.7212
27.8593 | -103.7224
27.863 | -103.7232
27.8667 | -103.7233
27.8704 | -103.7248
27.8741 | -103.7268
27.8778 | -103.7276
27.8815 | -103.7283
27.8852 | -103.731
27.8889 | -103.7336
27.8926 | -103.7332
27.8963 | -103.7297
27.9 | -103.7273
27.9037 | -103.7246
27.9074 | -103.72
27.9111 | -103.7167
27.9148 | -103.7144
27.9185 | -103.7111
27.9222 | -103.7062
27.9259 | -103.6983
27.9296 | -103.6913
27.9333 | -103.6856
27.937 | -103.6788
27.9407 | -103.6726
27.9444 | -103.6656
27.9481 | -103.6578
27.9519 | -103.6513
27.9556 | -103.6441
27.9593 | -103.6364
27.963 | -103.6288
27.9667 | -103.6184
27.9704 | -103.6074
27.9741 | -103.595
27.9778 | -103.5822
27.9815 | -103.5706
27.9852 | -103.5583
27.9889 | -103.5449
27.9926 | -103.5322
27.9963 | -103.5171
28 | -103.4989
28.0037 | -103.4815