## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 255 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 240
Hood | 02-92B/P | 202
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1013.4 | hPa |  | PRC PRO 01
Humidity | 54.6 | %/Hr
Internal temperature of calibration bench | 23.89 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1016 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.02 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3066 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 06/07/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 11/07/2018
Technical Manager | Anthony HUE | Approval |  | 16/07/2017
CEO | Robert GLEMOT | Approval |  | 16/07/2017
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4148 | 57.5368
0.012209 | -35.9556 | 38.4704
0.018313 | -34.9115 | 28.4894
0.024417 | -34.4586 | 22.1068
0.030521 | -34.3494 | 18.3234
0.036626 | -34.2082 | 15.2278
0.04273 | -34.1246 | 12.919
0.048834 | -34.0721 | 11.3333
0.054938 | -34.003 | 10.1291
0.061043 | -33.9297 | 9.001
0.067147 | -33.9193 | 8.1277
0.073251 | -33.9199 | 7.372
0.079355 | -33.9367 | 6.8112
0.08546 | -33.9532 | 6.3513
0.091564 | -33.9494 | 5.9731
0.097668 | -33.9163 | 5.712
0.10377 | -33.8701 | 5.3649
0.10988 | -33.8399 | 5.0026
0.11598 | -33.8308 | 4.6438
0.12209 | -33.8468 | 4.3644
0.12819 | -33.8784 | 4.1315
0.13429 | -33.9091 | 3.8943
0.1404 | -33.9259 | 3.6297
0.1465 | -33.9412 | 3.3784
0.15261 | -33.9484 | 3.1443
0.15871 | -33.9454 | 2.9502
0.16482 | -33.9316 | 2.7909
0.17092 | -33.9078 | 2.7245
0.17702 | -33.8895 | 2.6245
0.18313 | -33.8727 | 2.5394
0.18923 | -33.8605 | 2.4326
0.19534 | -33.8537 | 2.3291
0.20144 | -33.8526 | 2.2164
0.20754 | -33.8553 | 2.142
0.21365 | -33.8577 | 2.0667
0.21975 | -33.8619 | 2.0013
0.22586 | -33.8698 | 1.9357
0.23196 | -33.8783 | 1.8629
0.23807 | -33.8868 | 1.7854
0.24417 | -33.8888 | 1.6934
0.25027 | -33.8921 | 1.61
0.25638 | -33.8919 | 1.5191
0.26248 | -33.89 | 1.4332
0.26859 | -33.8875 | 1.3509
0.27469 | -33.8867 | 1.265
0.2808 | -33.8924 | 1.2012
0.2869 | -33.8967 | 1.1335
0.293 | -33.902 | 1.0767
0.29911 | -33.9081 | 1.0384
0.30521 | -33.9139 | 1.0022
0.31132 | -33.9179 | 0.98828
0.31742 | -33.9221 | 0.96604
0.32353 | -33.9254 | 0.94773
0.32963 | -33.9282 | 0.93225
0.33573 | -33.93 | 0.91849
0.34184 | -33.9324 | 0.90151
0.34794 | -33.9354 | 0.86614
0.35405 | -33.9385 | 0.83593
0.36015 | -33.9418 | 0.80228
0.36626 | -33.944 | 0.76519
0.37236 | -33.9463 | 0.71588
0.37846 | -33.9484 | 0.65949
0.38457 | -33.9473 | 0.60183
0.39067 | -33.9479 | 0.55244
0.39678 | -33.9463 | 0.49722
0.40288 | -33.944 | 0.44519
0.40899 | -33.9403 | 0.3942
0.41509 | -33.9359 | 0.34301
0.42119 | -33.931 | 0.29623
0.4273 | -33.9274 | 0.24795
0.4334 | -33.924 | 0.19895
0.43951 | -33.9211 | 0.15076
0.44561 | -33.9195 | 0.1056
0.45172 | -33.9209 | 0.068069
0.45782 | -33.9207 | 0.026522
0.46392 | -33.9218 | -0.01158
0.47003 | -33.9234 | -0.046413
0.47613 | -33.9254 | -0.078172
0.48224 | -33.9275 | -0.10838
0.48834 | -33.93 | -0.12969
0.49445 | -33.932 | -0.15421
0.50055 | -33.9345 | -0.17653
0.50665 | -33.9363 | -0.19463
0.51276 | -33.9375 | -0.21418
0.51886 | -33.9388 | -0.23123
0.52497 | -33.938 | -0.23778
0.53107 | -33.9384 | -0.24785
0.53717 | -33.9389 | -0.25888
0.54328 | -33.9387 | -0.26403
0.54938 | -33.9388 | -0.272
0.55549 | -33.9384 | -0.28093
0.56159 | -33.9382 | -0.29201
0.5677 | -33.9379 | -0.30256
0.5738 | -33.9374 | -0.31579
0.5799 | -33.9374 | -0.32794
0.58601 | -33.9378 | -0.34458
0.59211 | -33.9388 | -0.3675
0.59822 | -33.9394 | -0.38451
0.60432 | -33.9404 | -0.40623
0.61043 | -33.9413 | -0.42779
0.61653 | -33.9425 | -0.44963
0.62263 | -33.9435 | -0.47301
0.62874 | -33.9448 | -0.49368
0.63484 | -33.9458 | -0.51519
0.64095 | -33.947 | -0.53653
0.64705 | -33.9481 | -0.55682
0.65316 | -33.9491 | -0.5765
0.65926 | -33.9501 | -0.59903
0.66536 | -33.9496 | -0.61722
0.67147 | -33.9498 | -0.63367
0.67757 | -33.9503 | -0.65164
0.68368 | -33.9502 | -0.67143
0.68978 | -33.9503 | -0.68929
0.69589 | -33.9499 | -0.70727
0.70199 | -33.9492 | -0.72872
0.70809 | -33.9486 | -0.7486
0.7142 | -33.9482 | -0.7675
0.7203 | -33.9474 | -0.78681
0.72641 | -33.9466 | -0.80723
0.73251 | -33.9464 | -0.82773
0.73862 | -33.9464 | -0.85161
0.74472 | -33.9467 | -0.87066
0.75082 | -33.9463 | -0.8913
0.75693 | -33.9464 | -0.91435
0.76303 | -33.9464 | -0.93644
0.76914 | -33.9469 | -0.96298
0.77524 | -33.9471 | -0.98757
0.78135 | -33.9475 | -1.0106
0.78745 | -33.9481 | -1.035
0.79355 | -33.9485 | -1.0593
0.79966 | -33.9494 | -1.0847
0.80576 | -33.9506 | -1.1146
0.81187 | -33.9508 | -1.1436
0.81797 | -33.9514 | -1.1737
0.82408 | -33.952 | -1.2005
0.83018 | -33.9525 | -1.2312
0.83628 | -33.9534 | -1.2612
0.84239 | -33.9532 | -1.2887
0.84849 | -33.9532 | -1.3163
0.8546 | -33.9536 | -1.3435
0.8607 | -33.9537 | -1.3723
0.86681 | -33.9538 | -1.4002
0.87291 | -33.9533 | -1.4256
0.87901 | -33.9529 | -1.4476
0.88512 | -33.9527 | -1.4703
0.89122 | -33.9526 | -1.4932
0.89733 | -33.9524 | -1.5123
0.90343 | -33.9521 | -1.5332
0.90953 | -33.9523 | -1.5479
0.91564 | -33.9518 | -1.5636
0.92174 | -33.9514 | -1.5775
0.92785 | -33.951 | -1.5909
0.93395 | -33.951 | -1.6025
0.94006 | -33.951 | -1.6159
0.94616 | -33.9508 | -1.6205
0.95226 | -33.951 | -1.6318
0.95837 | -33.9511 | -1.6432
0.96447 | -33.9515 | -1.6534
0.97058 | -33.9515 | -1.6632
0.97668 | -33.9513 | -1.671
0.98279 | -33.9513 | -1.6745
0.98889 | -33.9516 | -1.6793
0.99499 | -33.9516 | -1.6877
1.0011 | -33.9519 | -1.6926
1.0072 | -33.9518 | -1.6964
1.0133 | -33.9518 | -1.7002
1.0194 | -33.9518 | -1.7056
1.0255 | -33.9515 | -1.7124
1.0316 | -33.9509 | -1.7165
1.0377 | -33.9504 | -1.7223
1.0438 | -33.9497 | -1.7268
1.0499 | -33.9482 | -1.7332
1.056 | -33.9471 | -1.738
1.0621 | -33.9459 | -1.7437
1.0682 | -33.9455 | -1.7486
1.0743 | -33.9443 | -1.7563
1.0805 | -33.9437 | -1.763
1.0866 | -33.9421 | -1.7736
1.0927 | -33.9408 | -1.7809
1.0988 | -33.94 | -1.7913
1.1049 | -33.9386 | -1.7988
1.111 | -33.9371 | -1.8097
1.1171 | -33.9357 | -1.8189
1.1232 | -33.9344 | -1.8328
1.1293 | -33.9333 | -1.8445
1.1354 | -33.9323 | -1.8566
1.1415 | -33.9311 | -1.8703
1.1476 | -33.9298 | -1.884
1.1537 | -33.9284 | -1.897
1.1598 | -33.9283 | -1.915
1.1659 | -33.9277 | -1.9268
1.172 | -33.927 | -1.9399
1.1781 | -33.9263 | -1.9537
1.1842 | -33.9257 | -1.9691
1.1903 | -33.9264 | -1.9841
1.1964 | -33.9263 | -1.9984
1.2025 | -33.9262 | -2.0134
1.2086 | -33.9266 | -2.029
1.2147 | -33.9269 | -2.0438
1.2209 | -33.9275 | -2.0572
1.227 | -33.9294 | -2.0722
1.2331 | -33.9303 | -2.0864
1.2392 | -33.9312 | -2.1021
1.2453 | -33.9324 | -2.1141
1.2514 | -33.9337 | -2.1262
1.2575 | -33.9352 | -2.1402
1.2636 | -33.9374 | -2.1517
1.2697 | -33.9391 | -2.1635
1.2758 | -33.9406 | -2.1775
1.2819 | -33.9424 | -2.191
1.288 | -33.944 | -2.2048
1.2941 | -33.946 | -2.2155
1.3002 | -33.9479 | -2.224
1.3063 | -33.9496 | -2.234
1.3124 | -33.9514 | -2.2441
1.3185 | -33.9533 | -2.2548
1.3246 | -33.9551 | -2.2658
1.3307 | -33.9571 | -2.2754
1.3368 | -33.9589 | -2.283
1.3429 | -33.9605 | -2.2919
1.349 | -33.9619 | -2.3005
1.3551 | -33.9636 | -2.3115
1.3613 | -33.9651 | -2.3194
1.3674 | -33.9664 | -2.3279
1.3735 | -33.9676 | -2.3361
1.3796 | -33.9689 | -2.3436
1.3857 | -33.9699 | -2.3507
1.3918 | -33.9709 | -2.3605
1.3979 | -33.9722 | -2.369
1.404 | -33.9729 | -2.379
1.4101 | -33.9738 | -2.3881
1.4162 | -33.9746 | -2.3966
1.4223 | -33.9754 | -2.4069
1.4284 | -33.9762 | -2.4152
1.4345 | -33.9768 | -2.4255
1.4406 | -33.9772 | -2.4364
1.4467 | -33.9779 | -2.4458
1.4528 | -33.9786 | -2.455
1.4589 | -33.9794 | -2.4654
1.465 | -33.9796 | -2.4732
1.4711 | -33.9806 | -2.4821
1.4772 | -33.9811 | -2.494
1.4833 | -33.9817 | -2.5061
1.4894 | -33.9824 | -2.5174
1.4955 | -33.9831 | -2.5288
1.5016 | -33.9838 | -2.5406
1.5078 | -33.9842 | -2.5548
1.5139 | -33.9849 | -2.5672
1.52 | -33.9854 | -2.5803
1.5261 | -33.9858 | -2.5917
1.5322 | -33.986 | -2.6041
1.5383 | -33.9863 | -2.6161
1.5444 | -33.9869 | -2.629
1.5505 | -33.9873 | -2.642
1.5566 | -33.9878 | -2.6545
1.5627 | -33.9884 | -2.6675
1.5688 | -33.9892 | -2.6804
1.5749 | -33.9894 | -2.6924
1.581 | -33.9897 | -2.7086
1.5871 | -33.9903 | -2.7238
1.5932 | -33.9907 | -2.7363
1.5993 | -33.991 | -2.7528
1.6054 | -33.9913 | -2.7673
1.6115 | -33.9917 | -2.7838
1.6176 | -33.992 | -2.7997
1.6237 | -33.9922 | -2.8137
1.6298 | -33.9925 | -2.8287
1.6359 | -33.9929 | -2.8429
1.642 | -33.9933 | -2.857
1.6482 | -33.9937 | -2.8736
1.6543 | -33.9941 | -2.8899
1.6604 | -33.9945 | -2.9042
1.6665 | -33.9946 | -2.9187
1.6726 | -33.9948 | -2.9345
1.6787 | -33.9953 | -2.9511
1.6848 | -33.9952 | -2.9672
1.6909 | -33.9955 | -2.983
1.697 | -33.9957 | -2.9989
1.7031 | -33.9957 | -3.0143
1.7092 | -33.9957 | -3.0289
1.7153 | -33.9959 | -3.0447
1.7214 | -33.9954 | -3.06
1.7275 | -33.9952 | -3.0755
1.7336 | -33.995 | -3.0914
1.7397 | -33.995 | -3.1064
1.7458 | -33.995 | -3.1228
1.7519 | -33.9948 | -3.1387
1.758 | -33.9947 | -3.1535
1.7641 | -33.9947 | -3.1677
1.7702 | -33.9945 | -3.1832
1.7763 | -33.9945 | -3.1968
1.7824 | -33.9943 | -3.2128
1.7885 | -33.994 | -3.2269
1.7947 | -33.9941 | -3.2417
1.8008 | -33.994 | -3.2544
1.8069 | -33.9938 | -3.268
1.813 | -33.9938 | -3.282
1.8191 | -33.9935 | -3.2962
1.8252 | -33.9933 | -3.3102
1.8313 | -33.9932 | -3.3245
1.8374 | -33.9932 | -3.338
1.8435 | -33.9932 | -3.3503
1.8496 | -33.9929 | -3.3647
1.8557 | -33.9928 | -3.3794
1.8618 | -33.9927 | -3.3912
1.8679 | -33.9927 | -3.4043
1.874 | -33.9927 | -3.4174
1.8801 | -33.9926 | -3.4304
1.8862 | -33.9927 | -3.4432
1.8923 | -33.9926 | -3.4554
1.8984 | -33.9923 | -3.4661
1.9045 | -33.9922 | -3.4772
1.9106 | -33.9919 | -3.4891
1.9167 | -33.9916 | -3.5004
1.9228 | -33.9915 | -3.5113
1.9289 | -33.9912 | -3.5202
1.9351 | -33.9913 | -3.5294
1.9412 | -33.9913 | -3.5398
1.9473 | -33.9914 | -3.5518
1.9534 | -33.9913 | -3.563
1.9595 | -33.9911 | -3.5745
1.9656 | -33.9906 | -3.5852
1.9717 | -33.9904 | -3.5957
1.9778 | -33.9905 | -3.6072
1.9839 | -33.9905 | -3.6187
1.99 | -33.9905 | -3.6286
1.9961 | -33.9903 | -3.639
2.0022 | -33.9902 | -3.649
2.0083 | -33.9899 | -3.659
2.0144 | -33.9899 | -3.6689
2.0205 | -33.9899 | -3.6807
2.0266 | -33.9901 | -3.6924
2.0327 | -33.9901 | -3.7035
2.0388 | -33.9904 | -3.7173
2.0449 | -33.9905 | -3.7276
2.051 | -33.9905 | -3.7376
2.0571 | -33.9904 | -3.7489
2.0632 | -33.9903 | -3.7595
2.0693 | -33.9903 | -3.7706
2.0754 | -33.9906 | -3.7822
2.0816 | -33.9907 | -3.7941
2.0877 | -33.9907 | -3.8052
2.0938 | -33.9909 | -3.8175
2.0999 | -33.9909 | -3.829
2.106 | -33.9911 | -3.8408
2.1121 | -33.9911 | -3.853
2.1182 | -33.9913 | -3.863
2.1243 | -33.9915 | -3.8735
2.1304 | -33.9916 | -3.8837
2.1365 | -33.9919 | -3.8951
2.1426 | -33.992 | -3.9082
2.1487 | -33.9925 | -3.9198
2.1548 | -33.9929 | -3.9326
2.1609 | -33.9935 | -3.9446
2.167 | -33.9939 | -3.957
2.1731 | -33.9942 | -3.9687
2.1792 | -33.9949 | -3.9797
2.1853 | -33.9954 | -3.9924
2.1914 | -33.9959 | -4.0045
2.1975 | -33.9964 | -4.0164
2.2036 | -33.9968 | -4.0294
2.2097 | -33.9972 | -4.0412
2.2158 | -33.9982 | -4.0514
2.222 | -33.9987 | -4.0614
2.2281 | -33.999 | -4.072
2.2342 | -33.9995 | -4.0849
2.2403 | -33.9999 | -4.0948
2.2464 | -34.0007 | -4.1057
2.2525 | -34.0011 | -4.1163
2.2586 | -34.0019 | -4.126
2.2647 | -34.0027 | -4.1353
2.2708 | -34.0033 | -4.1469
2.2769 | -34.0041 | -4.1582
2.283 | -34.0046 | -4.1675
2.2891 | -34.0051 | -4.1777
2.2952 | -34.0056 | -4.1874
2.3013 | -34.0062 | -4.1974
2.3074 | -34.0069 | -4.2068
2.3135 | -34.0073 | -4.218
2.3196 | -34.0079 | -4.2265
2.3257 | -34.0083 | -4.2369
2.3318 | -34.0086 | -4.2457
2.3379 | -34.009 | -4.2569
2.344 | -34.0093 | -4.2677
2.3501 | -34.0098 | -4.2783
2.3562 | -34.0101 | -4.2866
2.3623 | -34.0103 | -4.2952
2.3685 | -34.0105 | -4.304
2.3746 | -34.0108 | -4.3144
2.3807 | -34.0113 | -4.3232
2.3868 | -34.0114 | -4.3321
2.3929 | -34.0118 | -4.3414
2.399 | -34.012 | -4.3493
2.4051 | -34.0125 | -4.3588
2.4112 | -34.0128 | -4.3678
2.4173 | -34.0131 | -4.3767
2.4234 | -34.0131 | -4.3853
2.4295 | -34.0133 | -4.3938
2.4356 | -34.0134 | -4.4031
2.4417 | -34.0135 | -4.413
2.4478 | -34.0135 | -4.4225
2.4539 | -34.0138 | -4.4304
2.46 | -34.0135 | -4.4401
2.4661 | -34.0135 | -4.4493
2.4722 | -34.0138 | -4.4581
2.4783 | -34.0139 | -4.4665
2.4844 | -34.0139 | -4.4751
2.4905 | -34.0138 | -4.4843
2.4966 | -34.0135 | -4.4933
2.5027 | -34.0135 | -4.5014
2.5089 | -34.0134 | -4.51
2.515 | -34.0132 | -4.5193
2.5211 | -34.0131 | -4.5277
2.5272 | -34.0128 | -4.5369
2.5333 | -34.0127 | -4.5464
2.5394 | -34.0126 | -4.5555
2.5455 | -34.0124 | -4.565
2.5516 | -34.0124 | -4.5738
2.5577 | -34.0123 | -4.5838
2.5638 | -34.0121 | -4.5943
2.5699 | -34.0119 | -4.6046
2.576 | -34.0117 | -4.615
2.5821 | -34.0115 | -4.6251
2.5882 | -34.0113 | -4.6348
2.5943 | -34.0112 | -4.6447
2.6004 | -34.011 | -4.6548
2.6065 | -34.0108 | -4.6634
2.6126 | -34.0107 | -4.6727
2.6187 | -34.0106 | -4.6816
2.6248 | -34.0105 | -4.6907
2.6309 | -34.0103 | -4.702
2.637 | -34.0104 | -4.7133
2.6431 | -34.0103 | -4.7229
2.6492 | -34.0102 | -4.7331
2.6554 | -34.0102 | -4.7435
2.6615 | -34.01 | -4.7535
2.6676 | -34.01 | -4.7641
2.6737 | -34.0101 | -4.7749
2.6798 | -34.0101 | -4.785
2.6859 | -34.0101 | -4.7957
2.692 | -34.01 | -4.8055
2.6981 | -34.01 | -4.8167
2.7042 | -34.0101 | -4.8289
2.7103 | -34.01 | -4.8391
2.7164 | -34.0102 | -4.85
2.7225 | -34.0103 | -4.8614
2.7286 | -34.0104 | -4.8727
2.7347 | -34.0104 | -4.8829
2.7408 | -34.0107 | -4.8923
2.7469 | -34.0106 | -4.9037
2.753 | -34.011 | -4.9145
2.7591 | -34.0111 | -4.9265
2.7652 | -34.0111 | -4.9381
2.7713 | -34.0111 | -4.9481
2.7774 | -34.0115 | -4.9596
2.7835 | -34.0116 | -4.9703
2.7896 | -34.0118 | -4.9797
2.7958 | -34.0121 | -4.9901
2.8019 | -34.0126 | -5
2.808 | -34.0129 | -5.0119
2.8141 | -34.0133 | -5.0249
2.8202 | -34.0137 | -5.0371
2.8263 | -34.0141 | -5.048
2.8324 | -34.0144 | -5.0604
2.8385 | -34.0148 | -5.0725
2.8446 | -34.0153 | -5.0841
2.8507 | -34.0157 | -5.0962
2.8568 | -34.0161 | -5.1074
2.8629 | -34.0165 | -5.1187
2.869 | -34.017 | -5.1301
2.8751 | -34.0174 | -5.142
2.8812 | -34.018 | -5.1551
2.8873 | -34.0184 | -5.1664
2.8934 | -34.0187 | -5.1783
2.8995 | -34.019 | -5.1902
2.9056 | -34.0194 | -5.201
2.9117 | -34.0198 | -5.2125
2.9178 | -34.0205 | -5.2251
2.9239 | -34.0209 | -5.2384
2.93 | -34.0214 | -5.2518
2.9361 | -34.0219 | -5.2632
2.9423 | -34.0225 | -5.2752
2.9484 | -34.023 | -5.2886
2.9545 | -34.0237 | -5.3017
2.9606 | -34.0242 | -5.3149
2.9667 | -34.0247 | -5.3267
2.9728 | -34.025 | -5.3388
2.9789 | -34.0255 | -5.3514
2.985 | -34.0259 | -5.365
2.9911 | -34.0264 | -5.3774
2.9972 | -34.0269 | -5.39
3.0033 | -34.0275 | -5.4029
3.0094 | -34.0281 | -5.4159
3.0155 | -34.0286 | -5.4284
3.0216 | -34.0291 | -5.4415
3.0277 | -34.0295 | -5.455
3.0338 | -34.0301 | -5.4684
3.0399 | -34.0307 | -5.482
3.046 | -34.031 | -5.4949
3.0521 | -34.0316 | -5.5072
3.0582 | -34.0319 | -5.5208
3.0643 | -34.0324 | -5.5338
3.0704 | -34.0327 | -5.547
3.0765 | -34.0331 | -5.559
3.0827 | -34.0333 | -5.5719
3.0888 | -34.0337 | -5.5846
3.0949 | -34.034 | -5.5974
3.101 | -34.0344 | -5.6113
3.1071 | -34.0347 | -5.6252
3.1132 | -34.0352 | -5.639
3.1193 | -34.0356 | -5.6519
3.1254 | -34.036 | -5.6639
3.1315 | -34.0364 | -5.678
3.1376 | -34.0366 | -5.6907
3.1437 | -34.037 | -5.7041
3.1498 | -34.0373 | -5.7181
3.1559 | -34.0376 | -5.731
3.162 | -34.0377 | -5.7448
3.1681 | -34.0377 | -5.7588
3.1742 | -34.0381 | -5.7723
3.1803 | -34.0384 | -5.7854
3.1864 | -34.0388 | -5.7974
3.1925 | -34.0391 | -5.8109
3.1986 | -34.0392 | -5.8235
3.2047 | -34.0393 | -5.8372
3.2108 | -34.0395 | -5.8488
3.2169 | -34.0399 | -5.8623
3.223 | -34.0399 | -5.8755
3.2292 | -34.0398 | -5.8892
3.2353 | -34.0397 | -5.8986
3.2414 | -34.0399 | -5.9114
3.2475 | -34.0398 | -5.9223
3.2536 | -34.0401 | -5.9358
3.2597 | -34.0403 | -5.9469
3.2658 | -34.0405 | -5.9574
3.2719 | -34.0406 | -5.9683
3.278 | -34.0406 | -5.9803
3.2841 | -34.0408 | -5.9921
3.2902 | -34.0408 | -6.0045
3.2963 | -34.041 | -6.0161
3.3024 | -34.0408 | -6.0263
3.3085 | -34.0411 | -6.0386
3.3146 | -34.0409 | -6.0494
3.3207 | -34.0411 | -6.0614
3.3268 | -34.0412 | -6.0736
3.3329 | -34.0413 | -6.0832
3.339 | -34.0414 | -6.0928
3.3451 | -34.0414 | -6.1041
3.3512 | -34.0415 | -6.1134
3.3573 | -34.0416 | -6.1246
3.3634 | -34.0417 | -6.136
3.3696 | -34.0416 | -6.1478
3.3757 | -34.0416 | -6.1576
3.3818 | -34.0419 | -6.1692
3.3879 | -34.0418 | -6.1795
3.394 | -34.0417 | -6.1898
3.4001 | -34.0417 | -6.2006
3.4062 | -34.0417 | -6.2107
3.4123 | -34.0418 | -6.2204
3.4184 | -34.0417 | -6.2313
3.4245 | -34.0418 | -6.2417
3.4306 | -34.0417 | -6.2513
3.4367 | -34.0419 | -6.261
3.4428 | -34.0421 | -6.2702
3.4489 | -34.0421 | -6.2811
3.455 | -34.0419 | -6.2915
3.4611 | -34.0419 | -6.3025
3.4672 | -34.0419 | -6.3137
3.4733 | -34.042 | -6.3238
3.4794 | -34.0419 | -6.3342
3.4855 | -34.0422 | -6.3438
3.4916 | -34.0422 | -6.3534
3.4977 | -34.0422 | -6.3643
3.5038 | -34.0423 | -6.3743
3.5099 | -34.0423 | -6.3838
3.5161 | -34.0424 | -6.3933
3.5222 | -34.0425 | -6.4034
3.5283 | -34.0424 | -6.4133
3.5344 | -34.0424 | -6.424
3.5405 | -34.0425 | -6.4339
3.5466 | -34.0426 | -6.4436
3.5527 | -34.0427 | -6.4542
3.5588 | -34.0428 | -6.4639
3.5649 | -34.0428 | -6.474
3.571 | -34.0428 | -6.4842
3.5771 | -34.0429 | -6.4948
3.5832 | -34.043 | -6.5044
3.5893 | -34.0432 | -6.5142
3.5954 | -34.0433 | -6.5246
3.6015 | -34.0434 | -6.5349
3.6076 | -34.0436 | -6.5444
3.6137 | -34.0436 | -6.5543
3.6198 | -34.044 | -6.5632
3.6259 | -34.044 | -6.5739
3.632 | -34.0442 | -6.5844
3.6381 | -34.0444 | -6.5943
3.6442 | -34.0446 | -6.6046
3.6503 | -34.0448 | -6.6136
3.6565 | -34.0451 | -6.6253
3.6626 | -34.0451 | -6.6359
3.6687 | -34.0454 | -6.6465
3.6748 | -34.0457 | -6.6564
3.6809 | -34.046 | -6.6667
3.687 | -34.0462 | -6.6778
3.6931 | -34.0465 | -6.6877
3.6992 | -34.0467 | -6.6983
3.7053 | -34.0469 | -6.7083
3.7114 | -34.0471 | -6.7195
3.7175 | -34.0474 | -6.7293
3.7236 | -34.0476 | -6.738
3.7297 | -34.0477 | -6.7486
3.7358 | -34.0479 | -6.7602
3.7419 | -34.0481 | -6.7717
3.748 | -34.0482 | -6.7824
3.7541 | -34.0484 | -6.7933
3.7602 | -34.0487 | -6.8041
3.7663 | -34.0489 | -6.8142
3.7724 | -34.0492 | -6.8251
3.7785 | -34.0495 | -6.8354
3.7846 | -34.0497 | -6.8459
3.7907 | -34.0501 | -6.8568
3.7969 | -34.0503 | -6.8682
3.803 | -34.0506 | -6.8782
3.8091 | -34.0509 | -6.8884
3.8152 | -34.0512 | -6.8996
3.8213 | -34.0513 | -6.9111
3.8274 | -34.0517 | -6.9214
3.8335 | -34.0521 | -6.9328
3.8396 | -34.0525 | -6.9443
3.8457 | -34.0529 | -6.9544
3.8518 | -34.0534 | -6.9651
3.8579 | -34.0537 | -6.9756
3.864 | -34.054 | -6.9861
3.8701 | -34.0544 | -6.9969
3.8762 | -34.0547 | -7.0079
3.8823 | -34.055 | -7.0184
3.8884 | -34.0554 | -7.0287
3.8945 | -34.0558 | -7.0399
3.9006 | -34.0562 | -7.0513
3.9067 | -34.0565 | -7.063
3.9128 | -34.0569 | -7.0741
3.9189 | -34.0573 | -7.0846
3.925 | -34.0575 | -7.0954
3.9311 | -34.0579 | -7.1053
3.9372 | -34.0583 | -7.1167
3.9434 | -34.0586 | -7.1271
3.9495 | -34.059 | -7.1385
3.9556 | -34.0594 | -7.1497
3.9617 | -34.0597 | -7.1605
3.9678 | -34.0601 | -7.1711
3.9739 | -34.0605 | -7.1822
3.98 | -34.0608 | -7.194
3.9861 | -34.0611 | -7.2043
3.9922 | -34.0613 | -7.2159
3.9983 | -34.0617 | -7.2258
4.0044 | -34.062 | -7.2355
4.0105 | -34.0623 | -7.2464
4.0166 | -34.0628 | -7.2573
4.0227 | -34.0632 | -7.2675
4.0288 | -34.0634 | -7.2789
4.0349 | -34.0637 | -7.2886
4.041 | -34.064 | -7.2995
4.0471 | -34.0643 | -7.3097
4.0532 | -34.0647 | -7.3208
4.0593 | -34.065 | -7.3323
4.0654 | -34.0653 | -7.3429
4.0715 | -34.0656 | -7.3541
4.0776 | -34.066 | -7.3645
4.0838 | -34.0663 | -7.3758
4.0899 | -34.0666 | -7.3869
4.096 | -34.0669 | -7.3973
4.1021 | -34.0672 | -7.4081
4.1082 | -34.0676 | -7.4182
4.1143 | -34.0679 | -7.4283
4.1204 | -34.0683 | -7.439
4.1265 | -34.0687 | -7.4493
4.1326 | -34.069 | -7.4596
4.1387 | -34.0694 | -7.4707
4.1448 | -34.0698 | -7.4816
4.1509 | -34.07 | -7.4923
4.157 | -34.0702 | -7.5034
4.1631 | -34.0705 | -7.5146
4.1692 | -34.0709 | -7.5252
4.1753 | -34.0712 | -7.536
4.1814 | -34.0715 | -7.5457
4.1875 | -34.0718 | -7.557
4.1936 | -34.072 | -7.5683
4.1997 | -34.0723 | -7.5784
4.2058 | -34.0726 | -7.5899
4.2119 | -34.0728 | -7.5996
4.218 | -34.0731 | -7.6097
4.2241 | -34.0734 | -7.6193
4.2303 | -34.0738 | -7.6305
4.2364 | -34.0742 | -7.6418
4.2425 | -34.0744 | -7.6523
4.2486 | -34.0747 | -7.6634
4.2547 | -34.0748 | -7.6727
4.2608 | -34.0751 | -7.6826
4.2669 | -34.0753 | -7.6929
4.273 | -34.0756 | -7.7036
4.2791 | -34.0759 | -7.7144
4.2852 | -34.076 | -7.7246
4.2913 | -34.0762 | -7.7339
4.2974 | -34.0764 | -7.7448
4.3035 | -34.0767 | -7.7556
4.3096 | -34.0768 | -7.766
4.3157 | -34.077 | -7.7759
4.3218 | -34.0773 | -7.7859
4.3279 | -34.0775 | -7.7962
4.334 | -34.0777 | -7.806
4.3401 | -34.0779 | -7.8158
4.3462 | -34.0782 | -7.8266
4.3523 | -34.0785 | -7.8363
4.3584 | -34.0786 | -7.8447
4.3645 | -34.0788 | -7.8543
4.3707 | -34.079 | -7.8636
4.3768 | -34.0792 | -7.8739
4.3829 | -34.0794 | -7.8842
4.389 | -34.0797 | -7.8939
4.3951 | -34.0799 | -7.9034
4.4012 | -34.0801 | -7.9132
4.4073 | -34.0803 | -7.9233
4.4134 | -34.0805 | -7.9335
4.4195 | -34.0808 | -7.943
4.4256 | -34.081 | -7.9526
4.4317 | -34.0813 | -7.9621
4.4378 | -34.0814 | -7.9717
4.4439 | -34.0816 | -7.9811
4.45 | -34.0817 | -7.991
4.4561 | -34.0818 | -8.0008
4.4622 | -34.082 | -8.0103
4.4683 | -34.0821 | -8.0196
4.4744 | -34.0823 | -8.0297
4.4805 | -34.0826 | -8.0388
4.4866 | -34.0828 | -8.0479
4.4927 | -34.083 | -8.0581
4.4988 | -34.0832 | -8.0675
4.5049 | -34.0834 | -8.0766
4.511 | -34.0837 | -8.0857
4.5172 | -34.084 | -8.0955
4.5233 | -34.0842 | -8.106
4.5294 | -34.0844 | -8.1153
4.5355 | -34.0846 | -8.1246
4.5416 | -34.0849 | -8.1338
4.5477 | -34.0851 | -8.143
4.5538 | -34.0852 | -8.1517
4.5599 | -34.0856 | -8.1612
4.566 | -34.0857 | -8.1709
4.5721 | -34.086 | -8.18
4.5782 | -34.0863 | -8.1898
4.5843 | -34.0865 | -8.1994
4.5904 | -34.0868 | -8.2086
4.5965 | -34.0871 | -8.2173
4.6026 | -34.0873 | -8.2278
4.6087 | -34.0875 | -8.2367
4.6148 | -34.0878 | -8.2473
4.6209 | -34.0881 | -8.2558
4.627 | -34.0885 | -8.2658
4.6331 | -34.0885 | -8.2746
4.6392 | -34.0888 | -8.2838
4.6453 | -34.089 | -8.2932
4.6514 | -34.0893 | -8.3027
4.6576 | -34.0893 | -8.3132
4.6637 | -34.0896 | -8.3208
4.6698 | -34.0898 | -8.3294
4.6759 | -34.0899 | -8.3391
4.682 | -34.0901 | -8.3481
4.6881 | -34.0903 | -8.3568
4.6942 | -34.0906 | -8.3661
4.7003 | -34.0908 | -8.3755
4.7064 | -34.0911 | -8.3847
4.7125 | -34.0914 | -8.3941
4.7186 | -34.0915 | -8.4025
4.7247 | -34.0919 | -8.4121
4.7308 | -34.0922 | -8.4213
4.7369 | -34.0924 | -8.4309
4.743 | -34.0926 | -8.4411
4.7491 | -34.0928 | -8.4507
4.7552 | -34.0929 | -8.4599
4.7613 | -34.0931 | -8.4695
4.7674 | -34.0934 | -8.478
4.7735 | -34.0935 | -8.4881
4.7796 | -34.0938 | -8.4978
4.7857 | -34.0939 | -8.5074
4.7918 | -34.0941 | -8.516
4.7979 | -34.0945 | -8.525
4.8041 | -34.0946 | -8.5342
4.8102 | -34.0949 | -8.5432
4.8163 | -34.0951 | -8.5526
4.8224 | -34.0952 | -8.5612
4.8285 | -34.0956 | -8.5702
4.8346 | -34.0959 | -8.5797
4.8407 | -34.096 | -8.589
4.8468 | -34.0963 | -8.5987
4.8529 | -34.0965 | -8.6085
4.859 | -34.0967 | -8.6179
4.8651 | -34.097 | -8.6272
4.8712 | -34.0972 | -8.6366
4.8773 | -34.0974 | -8.6461
4.8834 | -34.0977 | -8.6558
4.8895 | -34.098 | -8.6659
4.8956 | -34.0981 | -8.6758
4.9017 | -34.0983 | -8.6851
4.9078 | -34.0985 | -8.695
4.9139 | -34.0988 | -8.7035
4.92 | -34.099 | -8.7125
4.9261 | -34.0993 | -8.7213
4.9322 | -34.0997 | -8.7305
4.9383 | -34.1 | -8.7395
4.9445 | -34.1003 | -8.7495
4.9506 | -34.1005 | -8.7583
4.9567 | -34.1007 | -8.7676
4.9628 | -34.101 | -8.777
4.9689 | -34.1012 | -8.7868
4.975 | -34.1014 | -8.7962
4.9811 | -34.1017 | -8.8059
4.9872 | -34.1019 | -8.8156
4.9933 | -34.1023 | -8.8253
4.9994 | -34.1026 | -8.8351
5.0055 | -34.103 | -8.8446
5.0116 | -34.1033 | -8.8548
5.0177 | -34.1037 | -8.8648
5.0238 | -34.104 | -8.8753
5.0299 | -34.1044 | -8.8862
5.036 | -34.1047 | -8.8957
5.0421 | -34.1049 | -8.9054
5.0482 | -34.1052 | -8.9156
5.0543 | -34.1055 | -8.925
5.0604 | -34.1058 | -8.9352
5.0665 | -34.1062 | -8.9449
5.0726 | -34.1065 | -8.9552
5.0787 | -34.1068 | -8.9654
5.0848 | -34.1071 | -8.9754
5.091 | -34.1074 | -8.9853
5.0971 | -34.1077 | -8.9953
5.1032 | -34.1081 | -9.0056
5.1093 | -34.1084 | -9.0156
5.1154 | -34.1086 | -9.0256
5.1215 | -34.1088 | -9.0356
5.1276 | -34.1092 | -9.0454
5.1337 | -34.1096 | -9.0551
5.1398 | -34.1098 | -9.0649
5.1459 | -34.1102 | -9.0753
5.152 | -34.1105 | -9.0863
5.1581 | -34.1107 | -9.0969
5.1642 | -34.111 | -9.1069
5.1703 | -34.1113 | -9.1169
5.1764 | -34.1116 | -9.1268
5.1825 | -34.1119 | -9.1371
5.1886 | -34.1122 | -9.1469
5.1947 | -34.1125 | -9.1577
5.2008 | -34.1129 | -9.1675
5.2069 | -34.1132 | -9.1787
5.213 | -34.1135 | -9.1883
5.2191 | -34.1139 | -9.1986
5.2252 | -34.1142 | -9.2089
5.2314 | -34.1144 | -9.2196
5.2375 | -34.1148 | -9.2296
5.2436 | -34.115 | -9.2398
5.2497 | -34.1154 | -9.2496
5.2558 | -34.1158 | -9.2606
5.2619 | -34.116 | -9.2714
5.268 | -34.1164 | -9.2819
5.2741 | -34.1167 | -9.2931
5.2802 | -34.1171 | -9.3039
5.2863 | -34.1174 | -9.3141
5.2924 | -34.1177 | -9.3252
5.2985 | -34.1181 | -9.3366
5.3046 | -34.1185 | -9.3478
5.3107 | -34.1189 | -9.3587
5.3168 | -34.1193 | -9.3691
5.3229 | -34.1196 | -9.3789
5.329 | -34.12 | -9.3889
5.3351 | -34.1204 | -9.399
5.3412 | -34.1208 | -9.4092
5.3473 | -34.1211 | -9.4201
5.3534 | -34.1215 | -9.431
5.3595 | -34.1218 | -9.4417
5.3656 | -34.1222 | -9.4526
5.3717 | -34.1225 | -9.4632
5.3779 | -34.1228 | -9.4741
5.384 | -34.1231 | -9.4849
5.3901 | -34.1234 | -9.4957
5.3962 | -34.1237 | -9.5065
5.4023 | -34.1241 | -9.5172
5.4084 | -34.1245 | -9.5276
5.4145 | -34.1247 | -9.5387
5.4206 | -34.125 | -9.5499
5.4267 | -34.1254 | -9.5613
5.4328 | -34.1257 | -9.5714
5.4389 | -34.1261 | -9.5816
5.445 | -34.1263 | -9.593
5.4511 | -34.1267 | -9.6033
5.4572 | -34.1271 | -9.6145
5.4633 | -34.1274 | -9.6249
5.4694 | -34.1277 | -9.6354
5.4755 | -34.128 | -9.6463
5.4816 | -34.1283 | -9.6569
5.4877 | -34.1287 | -9.6676
5.4938 | -34.129 | -9.6781
5.4999 | -34.1293 | -9.6888
5.506 | -34.1297 | -9.6995
5.5121 | -34.13 | -9.7105
5.5183 | -34.1303 | -9.7206
5.5244 | -34.1305 | -9.7314
5.5305 | -34.1309 | -9.7423
5.5366 | -34.1312 | -9.7528
5.5427 | -34.1315 | -9.764
5.5488 | -34.1318 | -9.775
5.5549 | -34.132 | -9.7857
5.561 | -34.1324 | -9.7969
5.5671 | -34.1328 | -9.8074
5.5732 | -34.133 | -9.8183
5.5793 | -34.1334 | -9.8298
5.5854 | -34.1338 | -9.8407
5.5915 | -34.1339 | -9.8516
5.5976 | -34.1342 | -9.8625
5.6037 | -34.1346 | -9.8731
5.6098 | -34.1348 | -9.8836
5.6159 | -34.1352 | -9.8945
5.622 | -34.1355 | -9.9049
5.6281 | -34.1359 | -9.9157
5.6342 | -34.1362 | -9.9266
5.6403 | -34.1364 | -9.9368
5.6464 | -34.1368 | -9.9475
5.6525 | -34.137 | -9.958
5.6586 | -34.1375 | -9.968
5.6648 | -34.1377 | -9.9792
5.6709 | -34.1381 | -9.9902
5.677 | -34.1384 | -10.0017
5.6831 | -34.1387 | -10.012
5.6892 | -34.139 | -10.0228
5.6953 | -34.1393 | -10.0334
5.7014 | -34.1396 | -10.0447
5.7075 | -34.1399 | -10.0555
5.7136 | -34.1402 | -10.066
5.7197 | -34.1405 | -10.0766
5.7258 | -34.1407 | -10.0872
5.7319 | -34.1411 | -10.0978
5.738 | -34.1414 | -10.1084
5.7441 | -34.1417 | -10.1191
5.7502 | -34.1421 | -10.1295
5.7563 | -34.1424 | -10.1398
5.7624 | -34.1427 | -10.1512
5.7685 | -34.1429 | -10.1617
5.7746 | -34.1431 | -10.1723
5.7807 | -34.1434 | -10.1826
5.7868 | -34.1437 | -10.1933
5.7929 | -34.1439 | -10.2038
5.799 | -34.1442 | -10.2144
5.8052 | -34.1445 | -10.2251
5.8113 | -34.1447 | -10.2357
5.8174 | -34.145 | -10.2464
5.8235 | -34.1453 | -10.2569
5.8296 | -34.1456 | -10.268
5.8357 | -34.1458 | -10.2786
5.8418 | -34.1461 | -10.2894
5.8479 | -34.1463 | -10.3003
5.854 | -34.1466 | -10.3112
5.8601 | -34.1469 | -10.3216
5.8662 | -34.1471 | -10.3325
5.8723 | -34.1474 | -10.3425
5.8784 | -34.1477 | -10.3531
5.8845 | -34.148 | -10.3639
5.8906 | -34.1482 | -10.3744
5.8967 | -34.1485 | -10.3846
5.9028 | -34.1489 | -10.3952
5.9089 | -34.1491 | -10.4052
5.915 | -34.1494 | -10.4158
5.9211 | -34.1497 | -10.4264
5.9272 | -34.1499 | -10.4368
5.9333 | -34.1502 | -10.4466
5.9394 | -34.1506 | -10.457
5.9455 | -34.1508 | -10.4671
5.9517 | -34.1512 | -10.4777
5.9578 | -34.1515 | -10.4885
5.9639 | -34.1518 | -10.4987
5.97 | -34.1521 | -10.5093
5.9761 | -34.1524 | -10.5195
5.9822 | -34.1527 | -10.5301
5.9883 | -34.153 | -10.5404
5.9944 | -34.1533 | -10.5509
6.0005 | -34.1537 | -10.5617
6.0066 | -34.1539 | -10.5717
6.0127 | -34.1543 | -10.5815
6.0188 | -34.1545 | -10.5924
6.0249 | -34.1548 | -10.6027
6.031 | -34.1551 | -10.613
6.0371 | -34.1555 | -10.6237
6.0432 | -34.1558 | -10.6341
6.0493 | -34.1561 | -10.6443
6.0554 | -34.1565 | -10.6547
6.0615 | -34.1568 | -10.6656
6.0676 | -34.157 | -10.6767
6.0737 | -34.1574 | -10.6873
6.0798 | -34.1576 | -10.6976
6.0859 | -34.158 | -10.708
6.0921 | -34.1583 | -10.7181
6.0982 | -34.1586 | -10.7287
6.1043 | -34.159 | -10.7389
6.1104 | -34.1593 | -10.7498
6.1165 | -34.1596 | -10.7594
6.1226 | -34.16 | -10.7693
6.1287 | -34.1603 | -10.7797
6.1348 | -34.1607 | -10.7901
6.1409 | -34.161 | -10.8004
6.147 | -34.1613 | -10.8107
6.1531 | -34.1617 | -10.821
6.1592 | -34.162 | -10.8314
6.1653 | -34.1623 | -10.8416
6.1714 | -34.1626 | -10.852
6.1775 | -34.163 | -10.8627
6.1836 | -34.1632 | -10.8728
6.1897 | -34.1636 | -10.8832
6.1958 | -34.1638 | -10.8938
6.2019 | -34.1642 | -10.9044
6.208 | -34.1645 | -10.9152
6.2141 | -34.1649 | -10.9257
6.2202 | -34.1652 | -10.9366
6.2263 | -34.1656 | -10.9467
6.2325 | -34.1659 | -10.9566
6.2386 | -34.1663 | -10.9666
6.2447 | -34.1665 | -10.9771
6.2508 | -34.1668 | -10.9873
6.2569 | -34.1672 | -10.9979
6.263 | -34.1676 | -11.0079
6.2691 | -34.1679 | -11.0181
6.2752 | -34.1683 | -11.0283
6.2813 | -34.1686 | -11.0386
6.2874 | -34.1689 | -11.049
6.2935 | -34.1693 | -11.0591
6.2996 | -34.1696 | -11.0694
6.3057 | -34.17 | -11.0797
6.3118 | -34.1703 | -11.0899
6.3179 | -34.1706 | -11.1001
6.324 | -34.171 | -11.1106
6.3301 | -34.1714 | -11.1207
6.3362 | -34.1717 | -11.131
6.3423 | -34.1721 | -11.1414
6.3484 | -34.1725 | -11.1519
6.3545 | -34.1728 | -11.1617
6.3606 | -34.1732 | -11.1716
6.3667 | -34.1736 | -11.1817
6.3728 | -34.174 | -11.192
6.379 | -34.1743 | -11.2022
6.3851 | -34.1746 | -11.2126
6.3912 | -34.175 | -11.2226
6.3973 | -34.1754 | -11.2327
6.4034 | -34.1757 | -11.2429
6.4095 | -34.176 | -11.2534
6.4156 | -34.1763 | -11.2632
6.4217 | -34.1767 | -11.2734
6.4278 | -34.177 | -11.2835
6.4339 | -34.1774 | -11.2938
6.44 | -34.1778 | -11.3037
6.4461 | -34.1781 | -11.3144
6.4522 | -34.1785 | -11.3241
6.4583 | -34.1788 | -11.3344
6.4644 | -34.1791 | -11.345
6.4705 | -34.1795 | -11.3554
6.4766 | -34.1798 | -11.3661
6.4827 | -34.1802 | -11.3768
6.4888 | -34.1805 | -11.3869
6.4949 | -34.1809 | -11.3967
6.501 | -34.1812 | -11.4073
6.5071 | -34.1816 | -11.4178
6.5132 | -34.182 | -11.4279
6.5194 | -34.1823 | -11.4381
6.5255 | -34.1827 | -11.448
6.5316 | -34.183 | -11.4584
6.5377 | -34.1834 | -11.4689
6.5438 | -34.1837 | -11.479
6.5499 | -34.1841 | -11.4892
6.556 | -34.1845 | -11.4993
6.5621 | -34.1848 | -11.5097
6.5682 | -34.1851 | -11.5197
6.5743 | -34.1855 | -11.5308
6.5804 | -34.1858 | -11.541
6.5865 | -34.186 | -11.5515
6.5926 | -34.1865 | -11.5617
6.5987 | -34.1868 | -11.5716
6.6048 | -34.1872 | -11.5824
6.6109 | -34.1876 | -11.5926
6.617 | -34.1879 | -11.6037
6.6231 | -34.1881 | -11.6142
6.6292 | -34.1886 | -11.6243
6.6353 | -34.1889 | -11.6355
6.6414 | -34.1893 | -11.6461
6.6475 | -34.1897 | -11.6565
6.6536 | -34.19 | -11.6664
6.6597 | -34.1905 | -11.6765
6.6659 | -34.1908 | -11.6867
6.672 | -34.1912 | -11.6969
6.6781 | -34.1915 | -11.707
6.6842 | -34.1919 | -11.7174
6.6903 | -34.1923 | -11.7276
6.6964 | -34.1926 | -11.7379
6.7025 | -34.1929 | -11.7482
6.7086 | -34.1933 | -11.7583
6.7147 | -34.1936 | -11.7687
6.7208 | -34.194 | -11.7795
6.7269 | -34.1943 | -11.7894
6.733 | -34.1946 | -11.7989
6.7391 | -34.195 | -11.8094
6.7452 | -34.1955 | -11.8191
6.7513 | -34.1958 | -11.8291
6.7574 | -34.1962 | -11.8396
6.7635 | -34.1966 | -11.8494
6.7696 | -34.197 | -11.8597
6.7757 | -34.1974 | -11.8703
6.7818 | -34.1978 | -11.8805
6.7879 | -34.1981 | -11.8909
6.794 | -34.1985 | -11.9011
6.8001 | -34.1988 | -11.9112
6.8063 | -34.1992 | -11.9216
6.8124 | -34.1995 | -11.932
6.8185 | -34.1999 | -11.9422
6.8246 | -34.2002 | -11.9525
6.8307 | -34.2006 | -11.9626
6.8368 | -34.201 | -11.9732
6.8429 | -34.2014 | -11.9833
6.849 | -34.2017 | -11.9936
6.8551 | -34.2021 | -12.0039
6.8612 | -34.2025 | -12.014
6.8673 | -34.2028 | -12.0241
6.8734 | -34.2032 | -12.035
6.8795 | -34.2036 | -12.0456
6.8856 | -34.2039 | -12.0553
6.8917 | -34.2043 | -12.0649
6.8978 | -34.2046 | -12.0751
6.9039 | -34.205 | -12.0851
6.91 | -34.2054 | -12.0954
6.9161 | -34.2058 | -12.1055
6.9222 | -34.2061 | -12.1158
6.9283 | -34.2064 | -12.1257
6.9344 | -34.2068 | -12.1357
6.9405 | -34.2072 | -12.1458
6.9466 | -34.2075 | -12.1562
6.9528 | -34.2079 | -12.1664
6.9589 | -34.2082 | -12.1765
6.965 | -34.2086 | -12.1867
6.9711 | -34.209 | -12.1969
6.9772 | -34.2094 | -12.2069
6.9833 | -34.2097 | -12.2173
6.9894 | -34.21 | -12.2279
6.9955 | -34.2104 | -12.2381
7.0016 | -34.2108 | -12.2487
7.0077 | -34.2112 | -12.2592
7.0138 | -34.2115 | -12.2696
7.0199 | -34.2118 | -12.2797
7.026 | -34.2121 | -12.2898
7.0321 | -34.2125 | -12.3
7.0382 | -34.2128 | -12.31
7.0443 | -34.2132 | -12.3204
7.0504 | -34.2135 | -12.3309
7.0565 | -34.2139 | -12.3414
7.0626 | -34.2142 | -12.3514
7.0687 | -34.2146 | -12.3613
7.0748 | -34.215 | -12.3712
7.0809 | -34.2153 | -12.3815
7.087 | -34.2156 | -12.3918
7.0932 | -34.216 | -12.4019
7.0993 | -34.2163 | -12.4123
7.1054 | -34.2166 | -12.4223
7.1115 | -34.2169 | -12.4324
7.1176 | -34.2172 | -12.4423
7.1237 | -34.2175 | -12.4525
7.1298 | -34.2178 | -12.4623
7.1359 | -34.2181 | -12.4721
7.142 | -34.2185 | -12.4824
7.1481 | -34.2188 | -12.4922
7.1542 | -34.2191 | -12.5022
7.1603 | -34.2194 | -12.5122
7.1664 | -34.2198 | -12.5224
7.1725 | -34.2202 | -12.5323
7.1786 | -34.2206 | -12.5425
7.1847 | -34.2209 | -12.5524
7.1908 | -34.2212 | -12.5627
7.1969 | -34.2215 | -12.5728
7.203 | -34.2219 | -12.5829
7.2091 | -34.2222 | -12.5927
7.2152 | -34.2225 | -12.6028
7.2213 | -34.2229 | -12.6127
7.2274 | -34.2232 | -12.6227
7.2335 | -34.2235 | -12.6328
7.2397 | -34.2239 | -12.6428
7.2458 | -34.2242 | -12.6525
7.2519 | -34.2245 | -12.6625
7.258 | -34.2249 | -12.6728
7.2641 | -34.2252 | -12.6827
7.2702 | -34.2255 | -12.693
7.2763 | -34.2259 | -12.7028
7.2824 | -34.2263 | -12.7127
7.2885 | -34.2266 | -12.7227
7.2946 | -34.227 | -12.7329
7.3007 | -34.2273 | -12.7429
7.3068 | -34.2276 | -12.7532
7.3129 | -34.2279 | -12.763
7.319 | -34.2283 | -12.7732
7.3251 | -34.2286 | -12.7828
7.3312 | -34.2289 | -12.7925
7.3373 | -34.2293 | -12.8023
7.3434 | -34.2297 | -12.812
7.3495 | -34.2301 | -12.8215
7.3556 | -34.2304 | -12.8319
7.3617 | -34.2307 | -12.8423
7.3678 | -34.2311 | -12.8524
7.3739 | -34.2314 | -12.862
7.3801 | -34.2317 | -12.8719
7.3862 | -34.232 | -12.8819
7.3923 | -34.2324 | -12.8917
7.3984 | -34.2327 | -12.9016
7.4045 | -34.2331 | -12.9114
7.4106 | -34.2334 | -12.9211
7.4167 | -34.2337 | -12.9313
7.4228 | -34.2341 | -12.9412
7.4289 | -34.2344 | -12.9508
7.435 | -34.2348 | -12.9605
7.4411 | -34.2352 | -12.9701
7.4472 | -34.2355 | -12.9802
7.4533 | -34.2359 | -12.9901
7.4594 | -34.2362 | -13.0003
7.4655 | -34.2366 | -13.0101
7.4716 | -34.2368 | -13.0203
7.4777 | -34.2372 | -13.0296
7.4838 | -34.2376 | -13.0397
7.4899 | -34.2379 | -13.0496
7.496 | -34.2383 | -13.0596
7.5021 | -34.2386 | -13.0694
7.5082 | -34.2389 | -13.0794
7.5143 | -34.2393 | -13.0891
7.5204 | -34.2396 | -13.0986
7.5266 | -34.2399 | -13.1086
7.5327 | -34.2403 | -13.1183
7.5388 | -34.2406 | -13.1282
7.5449 | -34.2409 | -13.1383
7.551 | -34.2412 | -13.1485
7.5571 | -34.2416 | -13.1582
7.5632 | -34.2419 | -13.168
7.5693 | -34.2423 | -13.1781
7.5754 | -34.2426 | -13.1879
7.5815 | -34.243 | -13.1974
7.5876 | -34.2433 | -13.2072
7.5937 | -34.2436 | -13.217
7.5998 | -34.2439 | -13.2269
7.6059 | -34.2442 | -13.2367
7.612 | -34.2446 | -13.2465
7.6181 | -34.2449 | -13.2563
7.6242 | -34.2453 | -13.2662
7.6303 | -34.2457 | -13.2759
7.6364 | -34.246 | -13.286
7.6425 | -34.2463 | -13.2962
7.6486 | -34.2466 | -13.3061
7.6547 | -34.247 | -13.3161
7.6608 | -34.2474 | -13.3259
7.667 | -34.2477 | -13.3352
7.6731 | -34.248 | -13.3451
7.6792 | -34.2483 | -13.3551
7.6853 | -34.2486 | -13.3648
7.6914 | -34.2488 | -13.3747
7.6975 | -34.2492 | -13.3847
7.7036 | -34.2495 | -13.3947
7.7097 | -34.2499 | -13.4044
7.7158 | -34.2502 | -13.4142
7.7219 | -34.2505 | -13.4238
7.728 | -34.2507 | -13.4338
7.7341 | -34.251 | -13.4438
7.7402 | -34.2514 | -13.4535
7.7463 | -34.2517 | -13.4635
7.7524 | -34.252 | -13.4737
7.7585 | -34.2523 | -13.4838
7.7646 | -34.2526 | -13.4933
7.7707 | -34.2529 | -13.503
7.7768 | -34.2531 | -13.5127
7.7829 | -34.2534 | -13.5225
7.789 | -34.2537 | -13.5322
7.7951 | -34.2541 | -13.5424
7.8012 | -34.2544 | -13.5526
7.8073 | -34.2547 | -13.5623
7.8135 | -34.255 | -13.5722
7.8196 | -34.2553 | -13.582
7.8257 | -34.2557 | -13.5923
7.8318 | -34.2559 | -13.6025
7.8379 | -34.2563 | -13.612
7.844 | -34.2565 | -13.6213
7.8501 | -34.2567 | -13.6315
7.8562 | -34.257 | -13.6419
7.8623 | -34.2574 | -13.6518
7.8684 | -34.2576 | -13.6613
7.8745 | -34.258 | -13.6709
7.8806 | -34.2582 | -13.6803
7.8867 | -34.2586 | -13.69
7.8928 | -34.2589 | -13.6997
7.8989 | -34.2593 | -13.7091
7.905 | -34.2596 | -13.7193
7.9111 | -34.2599 | -13.7291
7.9172 | -34.2602 | -13.7388
7.9233 | -34.2604 | -13.7487
7.9294 | -34.2608 | -13.7586
7.9355 | -34.2611 | -13.7679
7.9416 | -34.2614 | -13.7776
7.9477 | -34.2617 | -13.787
7.9539 | -34.262 | -13.7965
7.96 | -34.2623 | -13.8065
7.9661 | -34.2626 | -13.8164
7.9722 | -34.2629 | -13.8263
7.9783 | -34.2632 | -13.8361
7.9844 | -34.2636 | -13.8463
7.9905 | -34.2638 | -13.856
7.9966 | -34.2641 | -13.8657
8.0027 | -34.2644 | -13.8758
8.0088 | -34.2647 | -13.8853
8.0149 | -34.265 | -13.8948
8.021 | -34.2653 | -13.9045
8.0271 | -34.2656 | -13.9148
8.0332 | -34.2658 | -13.9246
8.0393 | -34.2662 | -13.9345
8.0454 | -34.2664 | -13.9444
8.0515 | -34.2668 | -13.9542
8.0576 | -34.267 | -13.9642
8.0637 | -34.2673 | -13.9735
8.0698 | -34.2676 | -13.9828
8.0759 | -34.2679 | -13.9923
8.082 | -34.2682 | -14.0021
8.0881 | -34.2685 | -14.0117
8.0942 | -34.2688 | -14.0218
8.1004 | -34.2691 | -14.032
8.1065 | -34.2694 | -14.0417
8.1126 | -34.2697 | -14.0513
8.1187 | -34.27 | -14.0613
8.1248 | -34.2703 | -14.0712
8.1309 | -34.2706 | -14.081
8.137 | -34.2709 | -14.0906
8.1431 | -34.2712 | -14.1005
8.1492 | -34.2715 | -14.1103
8.1553 | -34.2718 | -14.1202
8.1614 | -34.2721 | -14.1302
8.1675 | -34.2724 | -14.1398
8.1736 | -34.2727 | -14.1498
8.1797 | -34.273 | -14.1598
8.1858 | -34.2732 | -14.1695
8.1919 | -34.2735 | -14.1791
8.198 | -34.2738 | -14.1884
8.2041 | -34.2742 | -14.1985
8.2102 | -34.2745 | -14.2083
8.2163 | -34.2748 | -14.2179
8.2224 | -34.2751 | -14.2279
8.2285 | -34.2754 | -14.2375
8.2346 | -34.2757 | -14.2473
8.2408 | -34.276 | -14.2571
8.2469 | -34.2763 | -14.2667
8.253 | -34.2765 | -14.2763
8.2591 | -34.2769 | -14.2858
8.2652 | -34.2772 | -14.2953
8.2713 | -34.2776 | -14.3052
8.2774 | -34.2778 | -14.315
8.2835 | -34.2782 | -14.3251
8.2896 | -34.2785 | -14.3343
8.2957 | -34.2788 | -14.344
8.3018 | -34.2791 | -14.3537
8.3079 | -34.2794 | -14.3633
8.314 | -34.2797 | -14.373
8.3201 | -34.28 | -14.3827
8.3262 | -34.2803 | -14.3923
8.3323 | -34.2806 | -14.4019
8.3384 | -34.2809 | -14.412
8.3445 | -34.2812 | -14.4213
8.3506 | -34.2815 | -14.4312
8.3567 | -34.2818 | -14.4412
8.3628 | -34.2821 | -14.4508
8.3689 | -34.2824 | -14.4606
8.375 | -34.2827 | -14.4702
8.3812 | -34.283 | -14.4798
8.3873 | -34.2833 | -14.4896
8.3934 | -34.2836 | -14.4993
8.3995 | -34.2839 | -14.509
8.4056 | -34.2842 | -14.5188
8.4117 | -34.2845 | -14.5288
8.4178 | -34.2848 | -14.5381
8.4239 | -34.2851 | -14.5476
8.43 | -34.2855 | -14.5573
8.4361 | -34.2858 | -14.5675
8.4422 | -34.286 | -14.5771
8.4483 | -34.2863 | -14.5867
8.4544 | -34.2866 | -14.5966
8.4605 | -34.2869 | -14.6063
8.4666 | -34.2871 | -14.616
8.4727 | -34.2874 | -14.6256
8.4788 | -34.2877 | -14.635
8.4849 | -34.288 | -14.645
8.491 | -34.2883 | -14.6545
8.4971 | -34.2886 | -14.6644
8.5032 | -34.2889 | -14.6742
8.5093 | -34.2892 | -14.6837
8.5154 | -34.2895 | -14.6931
8.5215 | -34.2898 | -14.7028
8.5277 | -34.2901 | -14.7127
8.5338 | -34.2904 | -14.7221
8.5399 | -34.2907 | -14.732
8.546 | -34.291 | -14.7413
8.5521 | -34.2913 | -14.7512
8.5582 | -34.2916 | -14.7609
8.5643 | -34.2919 | -14.7703
8.5704 | -34.2922 | -14.7796
8.5765 | -34.2925 | -14.7892
8.5826 | -34.2929 | -14.7988
8.5887 | -34.2931 | -14.8083
8.5948 | -34.2934 | -14.8175
8.6009 | -34.2937 | -14.8271
8.607 | -34.294 | -14.8371
8.6131 | -34.2943 | -14.8469
8.6192 | -34.2946 | -14.8563
8.6253 | -34.2949 | -14.8661
8.6314 | -34.2952 | -14.8756
8.6375 | -34.2955 | -14.8855
8.6436 | -34.2959 | -14.8952
8.6497 | -34.2962 | -14.9048
8.6558 | -34.2965 | -14.9143
8.6619 | -34.2967 | -14.9241
8.6681 | -34.297 | -14.9336
8.6742 | -34.2973 | -14.9434
8.6803 | -34.2977 | -14.9534
8.6864 | -34.298 | -14.9633
8.6925 | -34.2983 | -14.9728
8.6986 | -34.2986 | -14.9827
8.7047 | -34.299 | -14.9924
8.7108 | -34.2992 | -15.0021
8.7169 | -34.2995 | -15.0116
8.723 | -34.2999 | -15.0213
8.7291 | -34.3002 | -15.0308
8.7352 | -34.3005 | -15.0406
8.7413 | -34.3008 | -15.0501
8.7474 | -34.3011 | -15.0595
8.7535 | -34.3015 | -15.0693
8.7596 | -34.3018 | -15.0793
8.7657 | -34.3021 | -15.0889
8.7718 | -34.3024 | -15.0981
8.7779 | -34.3027 | -15.1083
8.784 | -34.303 | -15.1178
8.7901 | -34.3033 | -15.1279
8.7962 | -34.3036 | -15.1377
8.8023 | -34.304 | -15.147
8.8084 | -34.3043 | -15.1568
8.8146 | -34.3046 | -15.1664
8.8207 | -34.3049 | -15.176
8.8268 | -34.3053 | -15.1855
8.8329 | -34.3056 | -15.1953
8.839 | -34.3059 | -15.205
8.8451 | -34.3062 | -15.2144
8.8512 | -34.3066 | -15.2242
8.8573 | -34.3069 | -15.2336
8.8634 | -34.3073 | -15.2433
8.8695 | -34.3076 | -15.253
8.8756 | -34.3079 | -15.2625
8.8817 | -34.3082 | -15.2724
8.8878 | -34.3084 | -15.2823
8.8939 | -34.3087 | -15.2918
8.9 | -34.309 | -15.3011
8.9061 | -34.3094 | -15.3108
8.9122 | -34.3097 | -15.3203
8.9183 | -34.31 | -15.3298
8.9244 | -34.3104 | -15.3392
8.9305 | -34.3107 | -15.3489
8.9366 | -34.3111 | -15.3584
8.9427 | -34.3114 | -15.3681
8.9488 | -34.3117 | -15.3778
8.955 | -34.312 | -15.3874
8.9611 | -34.3123 | -15.3972
8.9672 | -34.3127 | -15.4073
8.9733 | -34.313 | -15.4168
8.9794 | -34.3134 | -15.4266
8.9855 | -34.3137 | -15.4362
8.9916 | -34.314 | -15.4457
8.9977 | -34.3144 | -15.4552
9.0038 | -34.3147 | -15.465
9.0099 | -34.3151 | -15.4747
9.016 | -34.3155 | -15.4843
9.0221 | -34.3158 | -15.4942
9.0282 | -34.3161 | -15.5039
9.0343 | -34.3165 | -15.5138
9.0404 | -34.3169 | -15.5232
9.0465 | -34.3172 | -15.5333
9.0526 | -34.3175 | -15.5429
9.0587 | -34.3179 | -15.5525
9.0648 | -34.3183 | -15.562
9.0709 | -34.3186 | -15.5714
9.077 | -34.3189 | -15.5817
9.0831 | -34.3192 | -15.5915
9.0892 | -34.3195 | -15.601
9.0953 | -34.3198 | -15.6107
9.1015 | -34.3202 | -15.6204
9.1076 | -34.3206 | -15.6301
9.1137 | -34.3209 | -15.6395
9.1198 | -34.3213 | -15.6487
9.1259 | -34.3216 | -15.6585
9.132 | -34.322 | -15.6685
9.1381 | -34.3224 | -15.678
9.1442 | -34.3228 | -15.6873
9.1503 | -34.3231 | -15.6971
9.1564 | -34.3234 | -15.7067
9.1625 | -34.3237 | -15.7164
9.1686 | -34.3241 | -15.7259
9.1747 | -34.3245 | -15.7356
9.1808 | -34.3248 | -15.7456
9.1869 | -34.3252 | -15.7555
9.193 | -34.3255 | -15.7654
9.1991 | -34.3258 | -15.7747
9.2052 | -34.3262 | -15.7842
9.2113 | -34.3266 | -15.7941
9.2174 | -34.327 | -15.8044
9.2235 | -34.3273 | -15.8138
9.2296 | -34.3277 | -15.8236
9.2357 | -34.3281 | -15.8332
9.2419 | -34.3284 | -15.8428
9.248 | -34.3288 | -15.8527
9.2541 | -34.3291 | -15.8625
9.2602 | -34.3295 | -15.8722
9.2663 | -34.3298 | -15.8817
9.2724 | -34.3302 | -15.8914
9.2785 | -34.3305 | -15.9011
9.2846 | -34.3309 | -15.911
9.2907 | -34.3312 | -15.9206
9.2968 | -34.3315 | -15.9304
9.3029 | -34.3319 | -15.9403
9.309 | -34.3323 | -15.9499
9.3151 | -34.3326 | -15.9595
9.3212 | -34.333 | -15.9693
9.3273 | -34.3334 | -15.9789
9.3334 | -34.3338 | -15.9884
9.3395 | -34.3342 | -15.9981
9.3456 | -34.3346 | -16.0076
9.3517 | -34.335 | -16.017
9.3578 | -34.3353 | -16.0265
9.3639 | -34.3357 | -16.0363
9.37 | -34.3361 | -16.0461
9.3761 | -34.3364 | -16.0556
9.3822 | -34.3368 | -16.0652
9.3884 | -34.3372 | -16.075
9.3945 | -34.3376 | -16.0848
9.4006 | -34.3379 | -16.0946
9.4067 | -34.3383 | -16.1044
9.4128 | -34.3387 | -16.1141
9.4189 | -34.3391 | -16.1239
9.425 | -34.3395 | -16.1335
9.4311 | -34.3398 | -16.1432
9.4372 | -34.3402 | -16.1529
9.4433 | -34.3406 | -16.1627
9.4494 | -34.3409 | -16.1724
9.4555 | -34.3413 | -16.1819
9.4616 | -34.3417 | -16.1916
9.4677 | -34.3421 | -16.2012
9.4738 | -34.3425 | -16.2112
9.4799 | -34.3428 | -16.2208
9.486 | -34.3431 | -16.2309
9.4921 | -34.3436 | -16.241
9.4982 | -34.3439 | -16.2503
9.5043 | -34.3443 | -16.2598
9.5104 | -34.3447 | -16.2694
9.5165 | -34.3451 | -16.279
9.5226 | -34.3455 | -16.2889
9.5288 | -34.3459 | -16.2986
9.5349 | -34.3462 | -16.3083
9.541 | -34.3466 | -16.318
9.5471 | -34.347 | -16.3278
9.5532 | -34.3474 | -16.3375
9.5593 | -34.3478 | -16.3472
9.5654 | -34.3482 | -16.3571
9.5715 | -34.3485 | -16.3671
9.5776 | -34.3489 | -16.377
9.5837 | -34.3493 | -16.3865
9.5898 | -34.3496 | -16.3963
9.5959 | -34.35 | -16.4059
9.602 | -34.3503 | -16.4156
9.6081 | -34.3507 | -16.4254
9.6142 | -34.3511 | -16.435
9.6203 | -34.3514 | -16.4448
9.6264 | -34.3517 | -16.4542
9.6325 | -34.3521 | -16.464
9.6386 | -34.3525 | -16.4738
9.6447 | -34.3528 | -16.4833
9.6508 | -34.3532 | -16.4931
9.6569 | -34.3536 | -16.5028
9.663 | -34.354 | -16.5125
9.6691 | -34.3545 | -16.5223
9.6753 | -34.3549 | -16.5319
9.6814 | -34.3553 | -16.5414
9.6875 | -34.3557 | -16.5508
9.6936 | -34.3561 | -16.5604
9.6997 | -34.3565 | -16.5703
9.7058 | -34.3569 | -16.5801
9.7119 | -34.3573 | -16.5898
9.718 | -34.3576 | -16.5998
9.7241 | -34.358 | -16.6096
9.7302 | -34.3584 | -16.6194
9.7363 | -34.3588 | -16.6288
9.7424 | -34.3592 | -16.6385
9.7485 | -34.3595 | -16.648
9.7546 | -34.3599 | -16.6574
9.7607 | -34.3603 | -16.6674
9.7668 | -34.3607 | -16.6774
9.7729 | -34.3611 | -16.687
9.779 | -34.3615 | -16.6966
9.7851 | -34.3619 | -16.7064
9.7912 | -34.3623 | -16.7166
9.7973 | -34.3626 | -16.7264
9.8034 | -34.363 | -16.7364
9.8095 | -34.3634 | -16.7459
9.8157 | -34.3638 | -16.7557
9.8218 | -34.3641 | -16.765
9.8279 | -34.3645 | -16.7748
9.834 | -34.3649 | -16.7844
9.8401 | -34.3653 | -16.7941
9.8462 | -34.3657 | -16.8035
9.8523 | -34.3662 | -16.8131
9.8584 | -34.3665 | -16.8229
9.8645 | -34.3669 | -16.8326
9.8706 | -34.3673 | -16.8422
9.8767 | -34.3676 | -16.8515
9.8828 | -34.3681 | -16.8608
9.8889 | -34.3685 | -16.8707
9.895 | -34.3688 | -16.88
9.9011 | -34.3692 | -16.8893
9.9072 | -34.3696 | -16.8986
9.9133 | -34.37 | -16.9081
9.9194 | -34.3704 | -16.9176
9.9255 | -34.3708 | -16.9274
9.9316 | -34.3712 | -16.9367
9.9377 | -34.3716 | -16.9462
9.9438 | -34.372 | -16.9559
9.9499 | -34.3724 | -16.9656
9.956 | -34.3728 | -16.9751
9.9622 | -34.3732 | -16.9846
9.9683 | -34.3736 | -16.9944
9.9744 | -34.374 | -17.0039
9.9805 | -34.3744 | -17.0135
9.9866 | -34.3748 | -17.0228
9.9927 | -34.3752 | -17.0325
9.9988 | -34.3756 | -17.042
10.0049 | -34.376 | -17.0515
10.011 | -34.3764 | -17.0611
10.0171 | -34.3768 | -17.0706
10.0232 | -34.3772 | -17.08
10.0293 | -34.3777 | -17.0893
10.0354 | -34.3781 | -17.0986
10.0415 | -34.3785 | -17.1081
10.0476 | -34.3789 | -17.118
10.0537 | -34.3793 | -17.1279
10.0598 | -34.3797 | -17.1377
10.0659 | -34.38 | -17.1473
10.072 | -34.3805 | -17.1568
10.0781 | -34.3809 | -17.1664
10.0842 | -34.3813 | -17.1761
10.0903 | -34.3817 | -17.1859
10.0964 | -34.3822 | -17.1951
10.1026 | -34.3826 | -17.2046
10.1087 | -34.383 | -17.2142
10.1148 | -34.3834 | -17.2239
10.1209 | -34.3838 | -17.2333
10.127 | -34.3843 | -17.2427
10.1331 | -34.3847 | -17.2521
10.1392 | -34.3851 | -17.2616
10.1453 | -34.3855 | -17.2712
10.1514 | -34.3858 | -17.2804
10.1575 | -34.3862 | -17.2899
10.1636 | -34.3867 | -17.2988
10.1697 | -34.3871 | -17.3082
10.1758 | -34.3875 | -17.3176
10.1819 | -34.3879 | -17.3273
10.188 | -34.3883 | -17.3365
10.1941 | -34.3887 | -17.3462
10.2002 | -34.3891 | -17.3551
10.2063 | -34.3895 | -17.3649
10.2124 | -34.3899 | -17.3744
10.2185 | -34.3903 | -17.3838
10.2246 | -34.3907 | -17.3933
10.2307 | -34.391 | -17.4029
10.2368 | -34.3915 | -17.4116
10.2429 | -34.3919 | -17.4207
10.2491 | -34.3923 | -17.43
10.2552 | -34.3927 | -17.4398
10.2613 | -34.3932 | -17.4495
10.2674 | -34.3935 | -17.4591
10.2735 | -34.3939 | -17.4688
10.2796 | -34.3943 | -17.4784
10.2857 | -34.3947 | -17.4879
10.2918 | -34.3951 | -17.4975
10.2979 | -34.3955 | -17.5071
10.304 | -34.396 | -17.5166
10.3101 | -34.3964 | -17.5261
10.3162 | -34.3968 | -17.5355
10.3223 | -34.3972 | -17.5449
10.3284 | -34.3977 | -17.5544
10.3345 | -34.3981 | -17.5638
10.3406 | -34.3984 | -17.5731
10.3467 | -34.3989 | -17.5824
10.3528 | -34.3993 | -17.5916
10.3589 | -34.3996 | -17.6012
10.365 | -34.4 | -17.6108
10.3711 | -34.4004 | -17.6205
10.3772 | -34.4007 | -17.6299
10.3833 | -34.4011 | -17.6393
10.3895 | -34.4015 | -17.6485
10.3956 | -34.402 | -17.6577
10.4017 | -34.4023 | -17.667
10.4078 | -34.4027 | -17.6766
10.4139 | -34.4032 | -17.686
10.42 | -34.4036 | -17.6955
10.4261 | -34.4041 | -17.7048
10.4322 | -34.4045 | -17.7141
10.4383 | -34.4048 | -17.7235
10.4444 | -34.4052 | -17.7325
10.4505 | -34.4057 | -17.7419
10.4566 | -34.4061 | -17.7513
10.4627 | -34.4065 | -17.7605
10.4688 | -34.4069 | -17.7695
10.4749 | -34.4073 | -17.7789
10.481 | -34.4077 | -17.7882
10.4871 | -34.4081 | -17.7977
10.4932 | -34.4086 | -17.8069
10.4993 | -34.409 | -17.8162
10.5054 | -34.4094 | -17.8257
10.5115 | -34.4098 | -17.8352
10.5176 | -34.4103 | -17.8442
10.5237 | -34.4107 | -17.8536
10.5298 | -34.4111 | -17.8627
10.536 | -34.4116 | -17.8722
10.5421 | -34.412 | -17.8811
10.5482 | -34.4123 | -17.8908
10.5543 | -34.4128 | -17.8999
10.5604 | -34.4132 | -17.9088
10.5665 | -34.4137 | -17.918
10.5726 | -34.4141 | -17.9271
10.5787 | -34.4145 | -17.9366
10.5848 | -34.4149 | -17.9455
10.5909 | -34.4153 | -17.9549
10.597 | -34.4157 | -17.9641
10.6031 | -34.4161 | -17.9734
10.6092 | -34.4165 | -17.983
10.6153 | -34.4169 | -17.9925
10.6214 | -34.4174 | -18.0016
10.6275 | -34.4178 | -18.0107
10.6336 | -34.4183 | -18.0204
10.6397 | -34.4186 | -18.0296
10.6458 | -34.419 | -18.0387
10.6519 | -34.4194 | -18.048
10.658 | -34.4198 | -18.0569
10.6641 | -34.4202 | -18.0661
10.6702 | -34.4206 | -18.0757
10.6764 | -34.4211 | -18.0849
10.6825 | -34.4215 | -18.0943
10.6886 | -34.422 | -18.1038
10.6947 | -34.4224 | -18.1128
10.7008 | -34.4228 | -18.122
10.7069 | -34.4232 | -18.1313
10.713 | -34.4237 | -18.1404
10.7191 | -34.4241 | -18.1499
10.7252 | -34.4245 | -18.1589
10.7313 | -34.425 | -18.1681
10.7374 | -34.4254 | -18.177
10.7435 | -34.4258 | -18.1865
10.7496 | -34.4262 | -18.1955
10.7557 | -34.4266 | -18.2048
10.7618 | -34.4271 | -18.214
10.7679 | -34.4275 | -18.2233
10.774 | -34.4279 | -18.2325
10.7801 | -34.4283 | -18.2422
10.7862 | -34.4287 | -18.2513
10.7923 | -34.4292 | -18.2606
10.7984 | -34.4296 | -18.2696
10.8045 | -34.43 | -18.279
10.8106 | -34.4304 | -18.2882
10.8168 | -34.4309 | -18.2975
10.8229 | -34.4313 | -18.3067
10.829 | -34.4317 | -18.316
10.8351 | -34.4321 | -18.3253
10.8412 | -34.4325 | -18.3345
10.8473 | -34.4329 | -18.3435
10.8534 | -34.4333 | -18.3529
10.8595 | -34.4337 | -18.3619
10.8656 | -34.4342 | -18.3709
10.8717 | -34.4346 | -18.3796
10.8778 | -34.4351 | -18.3887
10.8839 | -34.4355 | -18.3977
10.89 | -34.4359 | -18.4068
10.8961 | -34.4363 | -18.416
10.9022 | -34.4368 | -18.425
10.9083 | -34.4372 | -18.4345
10.9144 | -34.4377 | -18.4436
10.9205 | -34.4381 | -18.4527
10.9266 | -34.4386 | -18.4623
10.9327 | -34.4391 | -18.4713
10.9388 | -34.4394 | -18.4805
10.9449 | -34.4398 | -18.4898
10.951 | -34.4402 | -18.499
10.9571 | -34.4407 | -18.5082
10.9633 | -34.4411 | -18.5178
10.9694 | -34.4415 | -18.5273
10.9755 | -34.442 | -18.5363
10.9816 | -34.4424 | -18.5455
10.9877 | -34.4429 | -18.5547
10.9938 | -34.4433 | -18.564
10.9999 | -34.4437 | -18.5731
11.006 | -34.4442 | -18.5821
11.0121 | -34.4446 | -18.5916
11.0182 | -34.4451 | -18.6004
11.0243 | -34.4455 | -18.6095
11.0304 | -34.4459 | -18.6186
11.0365 | -34.4463 | -18.6278
11.0426 | -34.4467 | -18.6368
11.0487 | -34.4472 | -18.646
11.0548 | -34.4476 | -18.655
11.0609 | -34.448 | -18.6645
11.067 | -34.4485 | -18.6738
11.0731 | -34.4489 | -18.6829
11.0792 | -34.4493 | -18.6918
11.0853 | -34.4497 | -18.7006
11.0914 | -34.4501 | -18.7097
11.0975 | -34.4506 | -18.7191
11.1037 | -34.451 | -18.728
11.1098 | -34.4514 | -18.7373
11.1159 | -34.4519 | -18.7466
11.122 | -34.4523 | -18.756
11.1281 | -34.4527 | -18.7649
11.1342 | -34.4531 | -18.7737
11.1403 | -34.4536 | -18.7831
11.1464 | -34.454 | -18.7922
11.1525 | -34.4545 | -18.8015
11.1586 | -34.4549 | -18.8111
11.1647 | -34.4554 | -18.8201
11.1708 | -34.4557 | -18.8291
11.1769 | -34.4561 | -18.8377
11.183 | -34.4566 | -18.8466
11.1891 | -34.4571 | -18.8561
11.1952 | -34.4575 | -18.8656
11.2013 | -34.4579 | -18.8747
11.2074 | -34.4583 | -18.8835
11.2135 | -34.4586 | -18.8926
11.2196 | -34.459 | -18.9015
11.2257 | -34.4594 | -18.9105
11.2318 | -34.4598 | -18.9197
11.2379 | -34.4601 | -18.9286
11.244 | -34.4605 | -18.9377
11.2502 | -34.461 | -18.9467
11.2563 | -34.4615 | -18.9556
11.2624 | -34.4618 | -18.9648
11.2685 | -34.4622 | -18.9736
11.2746 | -34.4626 | -18.9824
11.2807 | -34.463 | -18.992
11.2868 | -34.4634 | -19.0014
11.2929 | -34.4639 | -19.0101
11.299 | -34.4643 | -19.0194
11.3051 | -34.4647 | -19.0284
11.3112 | -34.4651 | -19.0377
11.3173 | -34.4656 | -19.0468
11.3234 | -34.466 | -19.056
11.3295 | -34.4664 | -19.0649
11.3356 | -34.4668 | -19.074
11.3417 | -34.4672 | -19.0832
11.3478 | -34.4676 | -19.0922
11.3539 | -34.468 | -19.1012
11.36 | -34.4684 | -19.1101
11.3661 | -34.4688 | -19.1191
11.3722 | -34.4693 | -19.128
11.3783 | -34.4697 | -19.137
11.3844 | -34.4701 | -19.1461
11.3906 | -34.4705 | -19.1555
11.3967 | -34.4709 | -19.1643
11.4028 | -34.4713 | -19.1736
11.4089 | -34.4717 | -19.1824
11.415 | -34.4721 | -19.1914
11.4211 | -34.4725 | -19.2002
11.4272 | -34.4729 | -19.2093
11.4333 | -34.4733 | -19.2183
11.4394 | -34.4737 | -19.2277
11.4455 | -34.4741 | -19.2368
11.4516 | -34.4745 | -19.2458
11.4577 | -34.4749 | -19.255
11.4638 | -34.4754 | -19.2637
11.4699 | -34.4758 | -19.2732
11.476 | -34.4762 | -19.2825
11.4821 | -34.4766 | -19.2913
11.4882 | -34.4771 | -19.3006
11.4943 | -34.4774 | -19.3094
11.5004 | -34.4779 | -19.3185
11.5065 | -34.4783 | -19.3276
11.5126 | -34.4788 | -19.3363
11.5187 | -34.4792 | -19.3454
11.5248 | -34.4796 | -19.3545
11.5309 | -34.48 | -19.3635
11.5371 | -34.4804 | -19.3725
11.5432 | -34.4808 | -19.3817
11.5493 | -34.4812 | -19.3908
11.5554 | -34.4816 | -19.3997
11.5615 | -34.482 | -19.4089
11.5676 | -34.4824 | -19.4181
11.5737 | -34.4829 | -19.4267
11.5798 | -34.4833 | -19.4354
11.5859 | -34.4838 | -19.4443
11.592 | -34.4841 | -19.4534
11.5981 | -34.4846 | -19.4623
11.6042 | -34.485 | -19.4715
11.6103 | -34.4854 | -19.4805
11.6164 | -34.4858 | -19.4893
11.6225 | -34.4862 | -19.4983
11.6286 | -34.4867 | -19.5075
11.6347 | -34.4871 | -19.5168
11.6408 | -34.4875 | -19.5257
11.6469 | -34.488 | -19.5348
11.653 | -34.4884 | -19.5436
11.6591 | -34.4888 | -19.5527
11.6652 | -34.4892 | -19.5617
11.6713 | -34.4896 | -19.5708
11.6775 | -34.4901 | -19.5796
11.6836 | -34.4905 | -19.5887
11.6897 | -34.4909 | -19.5975
11.6958 | -34.4913 | -19.6068
11.7019 | -34.4918 | -19.6159
11.708 | -34.4922 | -19.625
11.7141 | -34.4926 | -19.6339
11.7202 | -34.493 | -19.6427
11.7263 | -34.4934 | -19.6517
11.7324 | -34.4939 | -19.6609
11.7385 | -34.4942 | -19.6701
11.7446 | -34.4946 | -19.6792
11.7507 | -34.4951 | -19.688
11.7568 | -34.4955 | -19.6971
11.7629 | -34.4959 | -19.706
11.769 | -34.4963 | -19.7149
11.7751 | -34.4967 | -19.7237
11.7812 | -34.4971 | -19.7325
11.7873 | -34.4975 | -19.7418
11.7934 | -34.498 | -19.7509
11.7995 | -34.4984 | -19.7596
11.8056 | -34.4988 | -19.7689
11.8117 | -34.4992 | -19.778
11.8178 | -34.4996 | -19.7869
11.824 | -34.5001 | -19.7956
11.8301 | -34.5005 | -19.8042
11.8362 | -34.5009 | -19.8129
11.8423 | -34.5013 | -19.8219
11.8484 | -34.5017 | -19.8308
11.8545 | -34.5021 | -19.84
11.8606 | -34.5026 | -19.8489
11.8667 | -34.503 | -19.858
11.8728 | -34.5034 | -19.8669
11.8789 | -34.5038 | -19.8763
11.885 | -34.5041 | -19.8851
11.8911 | -34.5046 | -19.8943
11.8972 | -34.505 | -19.9029
11.9033 | -34.5054 | -19.9114
11.9094 | -34.5058 | -19.92
11.9155 | -34.5062 | -19.929
11.9216 | -34.5066 | -19.938
11.9277 | -34.507 | -19.9467
11.9338 | -34.5074 | -19.9558
11.9399 | -34.5078 | -19.965
11.946 | -34.5083 | -19.9738
11.9521 | -34.5087 | -19.9826
11.9582 | -34.5091 | -19.9915
11.9644 | -34.5096 | -20.0002
11.9705 | -34.51 | -20.0093
11.9766 | -34.5104 | -20.0182
11.9827 | -34.5108 | -20.027
11.9888 | -34.5112 | -20.0356
11.9949 | -34.5116 | -20.0445
12.001 | -34.512 | -20.0534
12.0071 | -34.5125 | -20.062
12.0132 | -34.5129 | -20.0706
12.0193 | -34.5133 | -20.0798
12.0254 | -34.5137 | -20.0884
12.0315 | -34.5142 | -20.0972
12.0376 | -34.5146 | -20.1059
12.0437 | -34.515 | -20.1146
12.0498 | -34.5154 | -20.1238
12.0559 | -34.5158 | -20.1327
12.062 | -34.5162 | -20.1416
12.0681 | -34.5166 | -20.1504
12.0742 | -34.5171 | -20.1593
12.0803 | -34.5175 | -20.1681
12.0864 | -34.5179 | -20.177
12.0925 | -34.5184 | -20.186
12.0986 | -34.5188 | -20.1945
12.1047 | -34.5192 | -20.2034
12.1109 | -34.5197 | -20.2124
12.117 | -34.5201 | -20.221
12.1231 | -34.5206 | -20.2298
12.1292 | -34.521 | -20.2388
12.1353 | -34.5214 | -20.2473
12.1414 | -34.5218 | -20.2562
12.1475 | -34.5222 | -20.2651
12.1536 | -34.5226 | -20.2741
12.1597 | -34.523 | -20.283
12.1658 | -34.5234 | -20.2919
12.1719 | -34.5239 | -20.3008
12.178 | -34.5244 | -20.3098
12.1841 | -34.5248 | -20.3188
12.1902 | -34.5252 | -20.3277
12.1963 | -34.5256 | -20.3365
12.2024 | -34.5261 | -20.3452
12.2085 | -34.5265 | -20.3541
12.2146 | -34.5269 | -20.3629
12.2207 | -34.5274 | -20.3717
12.2268 | -34.5278 | -20.3804
12.2329 | -34.5282 | -20.3892
12.239 | -34.5285 | -20.3978
12.2451 | -34.529 | -20.4064
12.2513 | -34.5295 | -20.4153
12.2574 | -34.5299 | -20.4243
12.2635 | -34.5303 | -20.4327
12.2696 | -34.5307 | -20.4414
12.2757 | -34.5311 | -20.4499
12.2818 | -34.5315 | -20.4584
12.2879 | -34.5319 | -20.4671
12.294 | -34.5323 | -20.4757
12.3001 | -34.5327 | -20.4847
12.3062 | -34.5332 | -20.4935
12.3123 | -34.5336 | -20.5022
12.3184 | -34.534 | -20.5108
12.3245 | -34.5345 | -20.5197
12.3306 | -34.5349 | -20.5282
12.3367 | -34.5353 | -20.5368
12.3428 | -34.5357 | -20.5453
12.3489 | -34.5362 | -20.5534
12.355 | -34.5366 | -20.5622
12.3611 | -34.537 | -20.571
12.3672 | -34.5374 | -20.5796
12.3733 | -34.5378 | -20.5887
12.3794 | -34.5382 | -20.597
12.3855 | -34.5386 | -20.6055
12.3916 | -34.539 | -20.6143
12.3978 | -34.5394 | -20.6233
12.4039 | -34.5398 | -20.632
12.41 | -34.5403 | -20.6409
12.4161 | -34.5407 | -20.6494
12.4222 | -34.5411 | -20.658
12.4283 | -34.5416 | -20.6667
12.4344 | -34.542 | -20.6753
12.4405 | -34.5424 | -20.684
12.4466 | -34.5428 | -20.6926
12.4527 | -34.5432 | -20.7013
12.4588 | -34.5437 | -20.7101
12.4649 | -34.5441 | -20.7191
12.471 | -34.5445 | -20.7283
12.4771 | -34.5449 | -20.7364
12.4832 | -34.5453 | -20.7451
12.4893 | -34.5458 | -20.7536
12.4954 | -34.5462 | -20.7621
12.5015 | -34.5467 | -20.7705
12.5076 | -34.5471 | -20.7792
12.5137 | -34.5475 | -20.7884
12.5198 | -34.5479 | -20.7967
12.5259 | -34.5483 | -20.8052
12.532 | -34.5488 | -20.814
12.5382 | -34.5492 | -20.8226
12.5443 | -34.5496 | -20.8313
12.5504 | -34.55 | -20.84
12.5565 | -34.5505 | -20.8486
12.5626 | -34.5509 | -20.8575
12.5687 | -34.5513 | -20.8663
12.5748 | -34.5518 | -20.8749
12.5809 | -34.5521 | -20.8832
12.587 | -34.5525 | -20.8922
12.5931 | -34.553 | -20.9009
12.5992 | -34.5534 | -20.9097
12.6053 | -34.5539 | -20.9182
12.6114 | -34.5543 | -20.9268
12.6175 | -34.5547 | -20.9356
12.6236 | -34.5551 | -20.9439
12.6297 | -34.5556 | -20.9524
12.6358 | -34.556 | -20.9608
12.6419 | -34.5565 | -20.9696
12.648 | -34.5569 | -20.9785
12.6541 | -34.5574 | -20.987
12.6602 | -34.5578 | -20.9952
12.6663 | -34.5582 | -21.0037
12.6724 | -34.5586 | -21.0124
12.6785 | -34.5591 | -21.0212
12.6847 | -34.5595 | -21.0299
12.6908 | -34.5599 | -21.0384
12.6969 | -34.5604 | -21.0469
12.703 | -34.5608 | -21.0556
12.7091 | -34.5612 | -21.0643
12.7152 | -34.5616 | -21.0729
12.7213 | -34.5621 | -21.0815
12.7274 | -34.5625 | -21.0903
12.7335 | -34.5629 | -21.0983
12.7396 | -34.5633 | -21.1065
12.7457 | -34.5638 | -21.1151
12.7518 | -34.5642 | -21.1236
12.7579 | -34.5646 | -21.1325
12.764 | -34.565 | -21.1411
12.7701 | -34.5653 | -21.1494
12.7762 | -34.5657 | -21.158
12.7823 | -34.5662 | -21.1667
12.7884 | -34.5665 | -21.1753
12.7945 | -34.567 | -21.1836
12.8006 | -34.5674 | -21.1916
12.8067 | -34.5678 | -21.2005
12.8128 | -34.5682 | -21.2087
12.8189 | -34.5686 | -21.2171
12.8251 | -34.569 | -21.2255
12.8312 | -34.5694 | -21.234
12.8373 | -34.5698 | -21.2426
12.8434 | -34.5703 | -21.2511
12.8495 | -34.5706 | -21.2596
12.8556 | -34.571 | -21.2681
12.8617 | -34.5714 | -21.2762
12.8678 | -34.5718 | -21.2848
12.8739 | -34.5722 | -21.2929
12.88 | -34.5726 | -21.3015
12.8861 | -34.573 | -21.3101
12.8922 | -34.5734 | -21.3188
12.8983 | -34.5739 | -21.327
12.9044 | -34.5743 | -21.3354
12.9105 | -34.5748 | -21.3437
12.9166 | -34.5752 | -21.3525
12.9227 | -34.5756 | -21.3615
12.9288 | -34.5761 | -21.37
12.9349 | -34.5765 | -21.3788
12.941 | -34.577 | -21.3873
12.9471 | -34.5775 | -21.3958
12.9532 | -34.5779 | -21.4044
12.9593 | -34.5783 | -21.4131
12.9654 | -34.5787 | -21.4218
12.9716 | -34.5791 | -21.4303
12.9777 | -34.5796 | -21.4384
12.9838 | -34.5799 | -21.4467
12.9899 | -34.5803 | -21.4553
12.996 | -34.5807 | -21.464
13.0021 | -34.5811 | -21.4724
13.0082 | -34.5815 | -21.4808
13.0143 | -34.5819 | -21.4889
13.0204 | -34.5823 | -21.4971
13.0265 | -34.5828 | -21.5053
13.0326 | -34.5831 | -21.5141
13.0387 | -34.5836 | -21.5227
13.0448 | -34.584 | -21.5311
13.0509 | -34.5844 | -21.5396
13.057 | -34.5849 | -21.5477
13.0631 | -34.5853 | -21.5559
13.0692 | -34.5858 | -21.5642
13.0753 | -34.5862 | -21.5725
13.0814 | -34.5866 | -21.5807
13.0875 | -34.587 | -21.5888
13.0936 | -34.5874 | -21.5972
13.0997 | -34.5878 | -21.6051
13.1058 | -34.5882 | -21.6131
13.112 | -34.5886 | -21.6214
13.1181 | -34.589 | -21.6298
13.1242 | -34.5894 | -21.6383
13.1303 | -34.5898 | -21.647
13.1364 | -34.5902 | -21.6553
13.1425 | -34.5906 | -21.6636
13.1486 | -34.591 | -21.6718
13.1547 | -34.5915 | -21.68
13.1608 | -34.5919 | -21.6884
13.1669 | -34.5924 | -21.6968
13.173 | -34.5928 | -21.7051
13.1791 | -34.5932 | -21.7136
13.1852 | -34.5936 | -21.7222
13.1913 | -34.594 | -21.7305
13.1974 | -34.5944 | -21.7388
13.2035 | -34.5948 | -21.7471
13.2096 | -34.5953 | -21.7554
13.2157 | -34.5957 | -21.7638
13.2218 | -34.5962 | -21.7722
13.2279 | -34.5966 | -21.7805
13.234 | -34.597 | -21.7889
13.2401 | -34.5974 | -21.7972
13.2462 | -34.5978 | -21.8056
13.2524 | -34.5982 | -21.8137
13.2585 | -34.5986 | -21.8219
13.2646 | -34.599 | -21.8304
13.2707 | -34.5994 | -21.8387
13.2768 | -34.5999 | -21.8473
13.2829 | -34.6003 | -21.8555
13.289 | -34.6007 | -21.8638
13.2951 | -34.6011 | -21.8719
13.3012 | -34.6015 | -21.8802
13.3073 | -34.6019 | -21.8888
13.3134 | -34.6023 | -21.897
13.3195 | -34.6028 | -21.9052
13.3256 | -34.6032 | -21.9136
13.3317 | -34.6036 | -21.9218
13.3378 | -34.604 | -21.9302
13.3439 | -34.6044 | -21.9385
13.35 | -34.6048 | -21.9467
13.3561 | -34.6052 | -21.955
13.3622 | -34.6056 | -21.9632
13.3683 | -34.6061 | -21.9714
13.3744 | -34.6065 | -21.9796
13.3805 | -34.6068 | -21.9881
13.3866 | -34.6072 | -21.9961
13.3927 | -34.6076 | -22.0044
13.3989 | -34.608 | -22.0124
13.405 | -34.6085 | -22.0205
13.4111 | -34.6089 | -22.0288
13.4172 | -34.6093 | -22.0367
13.4233 | -34.6097 | -22.0453
13.4294 | -34.6101 | -22.0533
13.4355 | -34.6106 | -22.0621
13.4416 | -34.6109 | -22.0705
13.4477 | -34.6113 | -22.0788
13.4538 | -34.6117 | -22.087
13.4599 | -34.6122 | -22.0953
13.466 | -34.6126 | -22.1038
13.4721 | -34.613 | -22.1122
13.4782 | -34.6135 | -22.1204
13.4843 | -34.6139 | -22.1286
13.4904 | -34.6143 | -22.1363
13.4965 | -34.6147 | -22.145
13.5026 | -34.6152 | -22.1532
13.5087 | -34.6156 | -22.1613
13.5148 | -34.616 | -22.1696
13.5209 | -34.6164 | -22.1777
13.527 | -34.6168 | -22.1863
13.5331 | -34.6172 | -22.1944
13.5393 | -34.6176 | -22.2028
13.5454 | -34.618 | -22.2109
13.5515 | -34.6185 | -22.2192
13.5576 | -34.6189 | -22.2277
13.5637 | -34.6193 | -22.2358
13.5698 | -34.6197 | -22.2443
13.5759 | -34.62 | -22.252
13.582 | -34.6204 | -22.2604
13.5881 | -34.6209 | -22.2685
13.5942 | -34.6213 | -22.2768
13.6003 | -34.6217 | -22.2849
13.6064 | -34.6221 | -22.2933
13.6125 | -34.6226 | -22.3019
13.6186 | -34.6231 | -22.3101
13.6247 | -34.6234 | -22.3184
13.6308 | -34.6238 | -22.3263
13.6369 | -34.6243 | -22.3345
13.643 | -34.6247 | -22.3427
13.6491 | -34.6252 | -22.3511
13.6552 | -34.6256 | -22.3592
13.6613 | -34.626 | -22.3675
13.6674 | -34.6264 | -22.3759
13.6735 | -34.6268 | -22.3844
13.6796 | -34.6272 | -22.3924
13.6858 | -34.6276 | -22.4007
13.6919 | -34.628 | -22.4091
13.698 | -34.6284 | -22.4175
13.7041 | -34.6288 | -22.426
13.7102 | -34.6292 | -22.4344
13.7163 | -34.6297 | -22.4427
13.7224 | -34.6301 | -22.4509
13.7285 | -34.6305 | -22.4593
13.7346 | -34.6309 | -22.4674
13.7407 | -34.6314 | -22.4758
13.7468 | -34.6318 | -22.484
13.7529 | -34.6321 | -22.492
13.759 | -34.6326 | -22.5005
13.7651 | -34.6329 | -22.5087
13.7712 | -34.6333 | -22.5171
13.7773 | -34.6338 | -22.5255
13.7834 | -34.6342 | -22.5336
13.7895 | -34.6347 | -22.5419
13.7956 | -34.6351 | -22.5503
13.8017 | -34.6355 | -22.5584
13.8078 | -34.636 | -22.5666
13.8139 | -34.6364 | -22.5746
13.82 | -34.6368 | -22.5825
13.8262 | -34.6373 | -22.591
13.8323 | -34.6377 | -22.5994
13.8384 | -34.6381 | -22.6079
13.8445 | -34.6386 | -22.6163
13.8506 | -34.639 | -22.6246
13.8567 | -34.6394 | -22.6331
13.8628 | -34.6398 | -22.6413
13.8689 | -34.6402 | -22.6496
13.875 | -34.6406 | -22.6577
13.8811 | -34.6411 | -22.666
13.8872 | -34.6415 | -22.6742
13.8933 | -34.642 | -22.6825
13.8994 | -34.6424 | -22.6906
13.9055 | -34.6428 | -22.6988
13.9116 | -34.6433 | -22.7068
13.9177 | -34.6437 | -22.7146
13.9238 | -34.6442 | -22.7231
13.9299 | -34.6447 | -22.7312
13.936 | -34.645 | -22.7393
13.9421 | -34.6454 | -22.7476
13.9482 | -34.6459 | -22.7559
13.9543 | -34.6463 | -22.7639
13.9604 | -34.6467 | -22.7722
13.9665 | -34.6472 | -22.7805
13.9727 | -34.6476 | -22.7887
13.9788 | -34.648 | -22.7971
13.9849 | -34.6484 | -22.8053
13.991 | -34.6488 | -22.814
13.9971 | -34.6492 | -22.8224
14.0032 | -34.6496 | -22.8308
14.0093 | -34.65 | -22.8389
14.0154 | -34.6504 | -22.8473
14.0215 | -34.6508 | -22.8553
14.0276 | -34.6513 | -22.8634
14.0337 | -34.6517 | -22.8718
14.0398 | -34.6521 | -22.8798
14.0459 | -34.6525 | -22.8882
14.052 | -34.6529 | -22.8966
14.0581 | -34.6533 | -22.9045
14.0642 | -34.6538 | -22.9129
14.0703 | -34.6542 | -22.9211
14.0764 | -34.6546 | -22.9296
14.0825 | -34.6551 | -22.9379
14.0886 | -34.6555 | -22.9458
14.0947 | -34.6559 | -22.9538
14.1008 | -34.6563 | -22.9623
14.1069 | -34.6567 | -22.9704
14.1131 | -34.6571 | -22.9789
14.1192 | -34.6575 | -22.9868
14.1253 | -34.658 | -22.9948
14.1314 | -34.6584 | -23.0029
14.1375 | -34.6589 | -23.0114
14.1436 | -34.6594 | -23.0196
14.1497 | -34.6598 | -23.0275
14.1558 | -34.6602 | -23.0359
14.1619 | -34.6607 | -23.0441
14.168 | -34.6611 | -23.052
14.1741 | -34.6616 | -23.06
14.1802 | -34.662 | -23.0682
14.1863 | -34.6624 | -23.0764
14.1924 | -34.6628 | -23.0846
14.1985 | -34.6633 | -23.0927
14.2046 | -34.6637 | -23.1008
14.2107 | -34.6642 | -23.1086
14.2168 | -34.6646 | -23.1172
14.2229 | -34.665 | -23.1255
14.229 | -34.6655 | -23.1337
14.2351 | -34.6659 | -23.1417
14.2412 | -34.6663 | -23.1496
14.2473 | -34.6668 | -23.1579
14.2534 | -34.6673 | -23.1658
14.2596 | -34.6677 | -23.174
14.2657 | -34.6681 | -23.1823
14.2718 | -34.6685 | -23.1906
14.2779 | -34.669 | -23.1985
14.284 | -34.6694 | -23.2068
14.2901 | -34.6699 | -23.2148
14.2962 | -34.6704 | -23.2227
14.3023 | -34.6708 | -23.231
14.3084 | -34.6712 | -23.2394
14.3145 | -34.6717 | -23.2475
14.3206 | -34.6721 | -23.2557
14.3267 | -34.6725 | -23.2637
14.3328 | -34.673 | -23.2718
14.3389 | -34.6734 | -23.2798
14.345 | -34.6738 | -23.2882
14.3511 | -34.6743 | -23.2961
14.3572 | -34.6747 | -23.3046
14.3633 | -34.6751 | -23.3126
14.3694 | -34.6755 | -23.3207
14.3755 | -34.6759 | -23.3285
14.3816 | -34.6764 | -23.3365
14.3877 | -34.6768 | -23.3445
14.3938 | -34.6773 | -23.3529
14.4 | -34.6777 | -23.3609
14.4061 | -34.6781 | -23.3691
14.4122 | -34.6785 | -23.3774
14.4183 | -34.6789 | -23.3856
14.4244 | -34.6794 | -23.3935
14.4305 | -34.6799 | -23.402
14.4366 | -34.6803 | -23.4102
14.4427 | -34.6808 | -23.4186
14.4488 | -34.6812 | -23.4264
14.4549 | -34.6816 | -23.4344
14.461 | -34.6821 | -23.4427
14.4671 | -34.6825 | -23.4509
14.4732 | -34.683 | -23.4586
14.4793 | -34.6834 | -23.4675
14.4854 | -34.6839 | -23.475
14.4915 | -34.6843 | -23.483
14.4976 | -34.6848 | -23.4911
14.5037 | -34.6853 | -23.4994
14.5098 | -34.6857 | -23.5073
14.5159 | -34.6862 | -23.5156
14.522 | -34.6866 | -23.5238
14.5281 | -34.6871 | -23.5317
14.5342 | -34.6875 | -23.54
14.5403 | -34.688 | -23.5481
14.5465 | -34.6885 | -23.5566
14.5526 | -34.6889 | -23.5649
14.5587 | -34.6893 | -23.5731
14.5648 | -34.6898 | -23.581
14.5709 | -34.6902 | -23.5893
14.577 | -34.6907 | -23.5975
14.5831 | -34.6911 | -23.6058
14.5892 | -34.6917 | -23.6136
14.5953 | -34.6921 | -23.6214
14.6014 | -34.6926 | -23.6295
14.6075 | -34.693 | -23.6371
14.6136 | -34.6934 | -23.6449
14.6197 | -34.6939 | -23.6532
14.6258 | -34.6943 | -23.6614
14.6319 | -34.6948 | -23.6695
14.638 | -34.6953 | -23.6776
14.6441 | -34.6957 | -23.6852
14.6502 | -34.6961 | -23.6931
14.6563 | -34.6966 | -23.7011
14.6624 | -34.697 | -23.7093
14.6685 | -34.6975 | -23.7172
14.6746 | -34.6979 | -23.7253
14.6807 | -34.6984 | -23.7333
14.6869 | -34.6988 | -23.7413
14.693 | -34.6993 | -23.7493
14.6991 | -34.6997 | -23.7572
14.7052 | -34.7002 | -23.7653
14.7113 | -34.7006 | -23.7736
14.7174 | -34.701 | -23.7816
14.7235 | -34.7015 | -23.7895
14.7296 | -34.7019 | -23.7975
14.7357 | -34.7024 | -23.8055
14.7418 | -34.7028 | -23.8133
14.7479 | -34.7033 | -23.8214
14.754 | -34.7038 | -23.8294
14.7601 | -34.7042 | -23.8375
14.7662 | -34.7047 | -23.8457
14.7723 | -34.7051 | -23.8535
14.7784 | -34.7056 | -23.8617
14.7845 | -34.706 | -23.8699
14.7906 | -34.7065 | -23.8779
14.7967 | -34.7069 | -23.8859
14.8028 | -34.7073 | -23.8937
14.8089 | -34.7078 | -23.9015
14.815 | -34.7082 | -23.9094
14.8211 | -34.7086 | -23.9173
14.8272 | -34.7091 | -23.9255
14.8334 | -34.7096 | -23.9334
14.8395 | -34.7101 | -23.9413
14.8456 | -34.7105 | -23.9493
14.8517 | -34.711 | -23.9574
14.8578 | -34.7114 | -23.9653
14.8639 | -34.7119 | -23.9735
14.87 | -34.7123 | -23.9815
14.8761 | -34.7128 | -23.9893
14.8822 | -34.7132 | -23.9973
14.8883 | -34.7137 | -24.0051
14.8944 | -34.7142 | -24.0132
14.9005 | -34.7147 | -24.021
14.9066 | -34.7151 | -24.029
14.9127 | -34.7156 | -24.037
14.9188 | -34.7161 | -24.0452
14.9249 | -34.7165 | -24.0531
14.931 | -34.717 | -24.0609
14.9371 | -34.7174 | -24.0686
14.9432 | -34.7179 | -24.0768
14.9493 | -34.7184 | -24.0849
14.9554 | -34.7188 | -24.0928
14.9615 | -34.7193 | -24.1008
14.9676 | -34.7198 | -24.1089
14.9738 | -34.7202 | -24.1168
14.9799 | -34.7206 | -24.1252
14.986 | -34.7211 | -24.1329
14.9921 | -34.7217 | -24.1408
14.9982 | -34.7221 | -24.1492
15.0043 | -34.7225 | -24.1572
15.0104 | -34.7229 | -24.1652
15.0165 | -34.7234 | -24.1733
15.0226 | -34.7239 | -24.181
15.0287 | -34.7243 | -24.1895
15.0348 | -34.7248 | -24.1976
15.0409 | -34.7253 | -24.2058
15.047 | -34.7258 | -24.2141
15.0531 | -34.7264 | -24.2225
15.0592 | -34.7268 | -24.2307
15.0653 | -34.7273 | -24.2388
15.0714 | -34.7277 | -24.2466
15.0775 | -34.7282 | -24.2544
15.0836 | -34.7287 | -24.2625
15.0897 | -34.7292 | -24.2706
15.0958 | -34.7297 | -24.2784
15.1019 | -34.7301 | -24.2865
15.108 | -34.7306 | -24.2947
15.1141 | -34.731 | -24.3026
15.1203 | -34.7315 | -24.3106
15.1264 | -34.7321 | -24.3188
15.1325 | -34.7325 | -24.3266
15.1386 | -34.733 | -24.3345
15.1447 | -34.7335 | -24.3426
15.1508 | -34.734 | -24.351
15.1569 | -34.7345 | -24.3593
15.163 | -34.735 | -24.367
15.1691 | -34.7355 | -24.3749
15.1752 | -34.736 | -24.3828
15.1813 | -34.7365 | -24.3908
15.1874 | -34.737 | -24.3989
15.1935 | -34.7375 | -24.4067
15.1996 | -34.738 | -24.415
15.2057 | -34.7385 | -24.4227
15.2118 | -34.7389 | -24.4307
15.2179 | -34.7394 | -24.4385
15.224 | -34.7398 | -24.4466
15.2301 | -34.7403 | -24.4543
15.2362 | -34.7408 | -24.4622
15.2423 | -34.7413 | -24.4704
15.2484 | -34.7418 | -24.4783
15.2545 | -34.7423 | -24.4863
15.2607 | -34.7428 | -24.4943
15.2668 | -34.7433 | -24.5022
15.2729 | -34.7438 | -24.51
15.279 | -34.7443 | -24.5181
15.2851 | -34.7448 | -24.5261
15.2912 | -34.7453 | -24.5339
15.2973 | -34.7457 | -24.5417
15.3034 | -34.7463 | -24.5494
15.3095 | -34.7467 | -24.557
15.3156 | -34.7472 | -24.5651
15.3217 | -34.7477 | -24.5727
15.3278 | -34.7482 | -24.5808
15.3339 | -34.7487 | -24.5884
15.34 | -34.7492 | -24.5963
15.3461 | -34.7496 | -24.6044
15.3522 | -34.7501 | -24.6122
15.3583 | -34.7506 | -24.6201
15.3644 | -34.7511 | -24.628
15.3705 | -34.7516 | -24.6356
15.3766 | -34.752 | -24.6438
15.3827 | -34.7525 | -24.6515
15.3888 | -34.753 | -24.6592
15.3949 | -34.7534 | -24.667
15.401 | -34.7539 | -24.675
15.4072 | -34.7545 | -24.6827
15.4133 | -34.755 | -24.6906
15.4194 | -34.7555 | -24.6985
15.4255 | -34.7559 | -24.7064
15.4316 | -34.7565 | -24.7145
15.4377 | -34.7569 | -24.7222
15.4438 | -34.7575 | -24.7298
15.4499 | -34.7579 | -24.7376
15.456 | -34.7583 | -24.7455
15.4621 | -34.7589 | -24.7534
15.4682 | -34.7594 | -24.7614
15.4743 | -34.7599 | -24.7693
15.4804 | -34.7604 | -24.7773
15.4865 | -34.7608 | -24.7852
15.4926 | -34.7613 | -24.7934
15.4987 | -34.7618 | -24.8013
15.5048 | -34.7623 | -24.809
15.5109 | -34.7629 | -24.8168
15.517 | -34.7634 | -24.8248
15.5231 | -34.7639 | -24.8327
15.5292 | -34.7643 | -24.8403
15.5353 | -34.7648 | -24.8483
15.5414 | -34.7653 | -24.8564
15.5476 | -34.7658 | -24.8644
15.5537 | -34.7663 | -24.8723
15.5598 | -34.7669 | -24.8799
15.5659 | -34.7673 | -24.8879
15.572 | -34.7678 | -24.8953
15.5781 | -34.7683 | -24.9032
15.5842 | -34.7688 | -24.9109
15.5903 | -34.7693 | -24.9186
15.5964 | -34.7699 | -24.9263
15.6025 | -34.7704 | -24.934
15.6086 | -34.7708 | -24.942
15.6147 | -34.7713 | -24.9499
15.6208 | -34.7718 | -24.9575
15.6269 | -34.7724 | -24.9653
15.633 | -34.7728 | -24.9732
15.6391 | -34.7733 | -24.9812
15.6452 | -34.7738 | -24.989
15.6513 | -34.7743 | -24.9969
15.6574 | -34.7748 | -25.0045
15.6635 | -34.7753 | -25.0125
15.6696 | -34.7758 | -25.0202
15.6757 | -34.7763 | -25.0283
15.6818 | -34.7768 | -25.0364
15.688 | -34.7774 | -25.0443
15.6941 | -34.7779 | -25.0521
15.7002 | -34.7784 | -25.06
15.7063 | -34.7788 | -25.068
15.7124 | -34.7793 | -25.076
15.7185 | -34.7798 | -25.0838
15.7246 | -34.7803 | -25.0914
15.7307 | -34.7808 | -25.0994
15.7368 | -34.7814 | -25.1076
15.7429 | -34.7819 | -25.1154
15.749 | -34.7824 | -25.1232
15.7551 | -34.7828 | -25.131
15.7612 | -34.7833 | -25.1383
15.7673 | -34.7838 | -25.146
15.7734 | -34.7843 | -25.1536
15.7795 | -34.7848 | -25.1618
15.7856 | -34.7853 | -25.1692
15.7917 | -34.7857 | -25.1769
15.7978 | -34.7863 | -25.1847
15.8039 | -34.7867 | -25.1923
15.81 | -34.7872 | -25.2
15.8161 | -34.7877 | -25.208
15.8222 | -34.7882 | -25.2157
15.8283 | -34.7888 | -25.2234
15.8345 | -34.7892 | -25.2312
15.8406 | -34.7898 | -25.2389
15.8467 | -34.7903 | -25.2469
15.8528 | -34.7908 | -25.2546
15.8589 | -34.7913 | -25.2624
15.865 | -34.7918 | -25.2701
15.8711 | -34.7923 | -25.2782
15.8772 | -34.7928 | -25.2861
15.8833 | -34.7933 | -25.2944
15.8894 | -34.7938 | -25.3022
15.8955 | -34.7943 | -25.3101
15.9016 | -34.7948 | -25.3179
15.9077 | -34.7953 | -25.3257
15.9138 | -34.7958 | -25.3335
15.9199 | -34.7963 | -25.3413
15.926 | -34.7968 | -25.3494
15.9321 | -34.7974 | -25.3572
15.9382 | -34.7979 | -25.365
15.9443 | -34.7984 | -25.373
15.9504 | -34.799 | -25.3808
15.9565 | -34.7995 | -25.3887
15.9626 | -34.8 | -25.3967
15.9687 | -34.8005 | -25.4043
15.9749 | -34.8011 | -25.4123
15.981 | -34.8016 | -25.4201
15.9871 | -34.8021 | -25.428
15.9932 | -34.8026 | -25.436
15.9993 | -34.8031 | -25.4437
16.0054 | -34.8036 | -25.4511
16.0115 | -34.8041 | -25.4588
16.0176 | -34.8046 | -25.4668
16.0237 | -34.8051 | -25.475
16.0298 | -34.8056 | -25.4823
16.0359 | -34.8061 | -25.49
16.042 | -34.8066 | -25.4979
16.0481 | -34.8071 | -25.5057
16.0542 | -34.8076 | -25.5137
16.0603 | -34.8082 | -25.5214
16.0664 | -34.8087 | -25.5292
16.0725 | -34.8092 | -25.5369
16.0786 | -34.8098 | -25.5447
16.0847 | -34.8102 | -25.5523
16.0908 | -34.8107 | -25.5601
16.0969 | -34.8113 | -25.5677
16.103 | -34.8117 | -25.5758
16.1091 | -34.8123 | -25.5835
16.1152 | -34.8127 | -25.5914
16.1214 | -34.8133 | -25.5992
16.1275 | -34.8138 | -25.607
16.1336 | -34.8143 | -25.6148
16.1397 | -34.8149 | -25.6225
16.1458 | -34.8154 | -25.6301
16.1519 | -34.8159 | -25.638
16.158 | -34.8164 | -25.646
16.1641 | -34.8169 | -25.6539
16.1702 | -34.8174 | -25.6618
16.1763 | -34.8179 | -25.6695
16.1824 | -34.8184 | -25.6772
16.1885 | -34.819 | -25.6848
16.1946 | -34.8195 | -25.6925
16.2007 | -34.82 | -25.7007
16.2068 | -34.8205 | -25.7084
16.2129 | -34.821 | -25.7163
16.219 | -34.8215 | -25.724
16.2251 | -34.8219 | -25.7319
16.2312 | -34.8224 | -25.7397
16.2373 | -34.8229 | -25.7474
16.2434 | -34.8234 | -25.7551
16.2495 | -34.824 | -25.7625
16.2556 | -34.8246 | -25.7704
16.2618 | -34.8251 | -25.7779
16.2679 | -34.8256 | -25.7857
16.274 | -34.8261 | -25.7936
16.2801 | -34.8266 | -25.8011
16.2862 | -34.8271 | -25.809
16.2923 | -34.8276 | -25.8169
16.2984 | -34.8281 | -25.8247
16.3045 | -34.8287 | -25.8324
16.3106 | -34.8292 | -25.84
16.3167 | -34.8297 | -25.8479
16.3228 | -34.8302 | -25.8559
16.3289 | -34.8307 | -25.8638
16.335 | -34.8312 | -25.8717
16.3411 | -34.8317 | -25.8793
16.3472 | -34.8322 | -25.8871
16.3533 | -34.8327 | -25.8949
16.3594 | -34.8332 | -25.9023
16.3655 | -34.8337 | -25.9103
16.3716 | -34.8341 | -25.9177
16.3777 | -34.8347 | -25.9255
16.3838 | -34.8352 | -25.9331
16.3899 | -34.8356 | -25.9408
16.396 | -34.8362 | -25.9486
16.4021 | -34.8367 | -25.9563
16.4083 | -34.8372 | -25.9638
16.4144 | -34.8377 | -25.9715
16.4205 | -34.8383 | -25.9795
16.4266 | -34.8387 | -25.987
16.4327 | -34.8393 | -25.9946
16.4388 | -34.8397 | -26.0024
16.4449 | -34.8402 | -26.0101
16.451 | -34.8407 | -26.0179
16.4571 | -34.8412 | -26.0259
16.4632 | -34.8417 | -26.0336
16.4693 | -34.8422 | -26.0414
16.4754 | -34.8427 | -26.0492
16.4815 | -34.8432 | -26.0571
16.4876 | -34.8437 | -26.0649
16.4937 | -34.8442 | -26.0725
16.4998 | -34.8447 | -26.0802
16.5059 | -34.8453 | -26.0878
16.512 | -34.8458 | -26.0959
16.5181 | -34.8463 | -26.1038
16.5242 | -34.8468 | -26.1116
16.5303 | -34.8473 | -26.1192
16.5364 | -34.8478 | -26.1268
16.5425 | -34.8482 | -26.1347
16.5487 | -34.8487 | -26.1424
16.5548 | -34.8492 | -26.1503
16.5609 | -34.8497 | -26.1582
16.567 | -34.8502 | -26.166
16.5731 | -34.8507 | -26.1739
16.5792 | -34.8512 | -26.1817
16.5853 | -34.8518 | -26.1891
16.5914 | -34.8523 | -26.1971
16.5975 | -34.8528 | -26.2049
16.6036 | -34.8533 | -26.2127
16.6097 | -34.8538 | -26.2204
16.6158 | -34.8542 | -26.2283
16.6219 | -34.8548 | -26.2358
16.628 | -34.8552 | -26.2436
16.6341 | -34.8557 | -26.2511
16.6402 | -34.8562 | -26.2584
16.6463 | -34.8567 | -26.2661
16.6524 | -34.8572 | -26.2739
16.6585 | -34.8577 | -26.2817
16.6646 | -34.8582 | -26.2895
16.6707 | -34.8587 | -26.2971
16.6768 | -34.8592 | -26.3048
16.6829 | -34.8597 | -26.3125
16.689 | -34.8602 | -26.3203
16.6952 | -34.8608 | -26.3282
16.7013 | -34.8613 | -26.3361
16.7074 | -34.8618 | -26.3439
16.7135 | -34.8623 | -26.3516
16.7196 | -34.8628 | -26.3593
16.7257 | -34.8633 | -26.3669
16.7318 | -34.8638 | -26.3746
16.7379 | -34.8643 | -26.3825
16.744 | -34.8648 | -26.3901
16.7501 | -34.8653 | -26.3978
16.7562 | -34.8658 | -26.4054
16.7623 | -34.8663 | -26.413
16.7684 | -34.8669 | -26.4208
16.7745 | -34.8674 | -26.4283
16.7806 | -34.8679 | -26.4361
16.7867 | -34.8684 | -26.4438
16.7928 | -34.8688 | -26.4514
16.7989 | -34.8693 | -26.459
16.805 | -34.8698 | -26.4668
16.8111 | -34.8703 | -26.4744
16.8172 | -34.8708 | -26.4824
16.8233 | -34.8713 | -26.4903
16.8294 | -34.8718 | -26.4982
16.8356 | -34.8723 | -26.5062
16.8417 | -34.8728 | -26.5139
16.8478 | -34.8733 | -26.5214
16.8539 | -34.8738 | -26.5291
16.86 | -34.8743 | -26.5368
16.8661 | -34.8747 | -26.5444
16.8722 | -34.8753 | -26.5523
16.8783 | -34.8758 | -26.5597
16.8844 | -34.8762 | -26.5673
16.8905 | -34.8768 | -26.5752
16.8966 | -34.8772 | -26.583
16.9027 | -34.8778 | -26.5905
16.9088 | -34.8783 | -26.5986
16.9149 | -34.8788 | -26.6066
16.921 | -34.8793 | -26.6143
16.9271 | -34.8798 | -26.6223
16.9332 | -34.8803 | -26.63
16.9393 | -34.8808 | -26.6379
16.9454 | -34.8813 | -26.6454
16.9515 | -34.8818 | -26.6533
16.9576 | -34.8823 | -26.661
16.9637 | -34.8828 | -26.6688
16.9698 | -34.8833 | -26.6763
16.9759 | -34.8838 | -26.6839
16.9821 | -34.8843 | -26.6918
16.9882 | -34.8848 | -26.6996
16.9943 | -34.8853 | -26.7074
17.0004 | -34.8858 | -26.715
17.0065 | -34.8863 | -26.7228
17.0126 | -34.8868 | -26.7304
17.0187 | -34.8873 | -26.7382
17.0248 | -34.8878 | -26.7456
17.0309 | -34.8883 | -26.7532
17.037 | -34.8887 | -26.761
17.0431 | -34.8893 | -26.7685
17.0492 | -34.8897 | -26.7762
17.0553 | -34.8902 | -26.7839
17.0614 | -34.8907 | -26.7915
17.0675 | -34.8912 | -26.7994
17.0736 | -34.8917 | -26.8073
17.0797 | -34.8922 | -26.8145
17.0858 | -34.8927 | -26.8225
17.0919 | -34.8931 | -26.8304
17.098 | -34.8936 | -26.8378
17.1041 | -34.8941 | -26.8454
17.1102 | -34.8946 | -26.8531
17.1163 | -34.8951 | -26.8609
17.1225 | -34.8956 | -26.8684
17.1286 | -34.8961 | -26.876
17.1347 | -34.8966 | -26.8835
17.1408 | -34.8971 | -26.8912
17.1469 | -34.8976 | -26.8987
17.153 | -34.8981 | -26.9065
17.1591 | -34.8986 | -26.9139
17.1652 | -34.899 | -26.9214
17.1713 | -34.8995 | -26.929
17.1774 | -34.9 | -26.9367
17.1835 | -34.9005 | -26.9445
17.1896 | -34.901 | -26.952
17.1957 | -34.9015 | -26.9599
17.2018 | -34.902 | -26.9676
17.2079 | -34.9024 | -26.9752
17.214 | -34.9029 | -26.9828
17.2201 | -34.9033 | -26.9905
17.2262 | -34.9038 | -26.9982
17.2323 | -34.9043 | -27.0058
17.2384 | -34.9048 | -27.0134
17.2445 | -34.9053 | -27.021
17.2506 | -34.9057 | -27.0289
17.2567 | -34.9062 | -27.0366
17.2628 | -34.9067 | -27.0445
17.269 | -34.9072 | -27.0521
17.2751 | -34.9077 | -27.0597
17.2812 | -34.9082 | -27.0672
17.2873 | -34.9086 | -27.0749
17.2934 | -34.9091 | -27.0826
17.2995 | -34.9096 | -27.0901
17.3056 | -34.91 | -27.0977
17.3117 | -34.9105 | -27.1053
17.3178 | -34.911 | -27.1128
17.3239 | -34.9114 | -27.1206
17.33 | -34.9119 | -27.128
17.3361 | -34.9124 | -27.1356
17.3422 | -34.9129 | -27.1432
17.3483 | -34.9133 | -27.1508
17.3544 | -34.9138 | -27.1583
17.3605 | -34.9143 | -27.166
17.3666 | -34.9148 | -27.1738
17.3727 | -34.9153 | -27.1813
17.3788 | -34.9158 | -27.1888
17.3849 | -34.9163 | -27.1964
17.391 | -34.9168 | -27.2039
17.3971 | -34.9172 | -27.2114
17.4032 | -34.9177 | -27.2194
17.4094 | -34.9182 | -27.2269
17.4155 | -34.9187 | -27.2344
17.4216 | -34.9191 | -27.2421
17.4277 | -34.9196 | -27.2499
17.4338 | -34.9201 | -27.2574
17.4399 | -34.9206 | -27.2652
17.446 | -34.921 | -27.2727
17.4521 | -34.9216 | -27.2802
17.4582 | -34.922 | -27.2877
17.4643 | -34.9225 | -27.2953
17.4704 | -34.923 | -27.3028
17.4765 | -34.9234 | -27.3104
17.4826 | -34.9239 | -27.318
17.4887 | -34.9244 | -27.3258
17.4948 | -34.9249 | -27.3332
17.5009 | -34.9253 | -27.3406
17.507 | -34.9258 | -27.3484
17.5131 | -34.9263 | -27.3559
17.5192 | -34.9268 | -27.3635
17.5253 | -34.9272 | -27.3713
17.5314 | -34.9277 | -27.3788
17.5375 | -34.9281 | -27.3865
17.5436 | -34.9286 | -27.3942
17.5497 | -34.9291 | -27.402
17.5559 | -34.9295 | -27.4092
17.562 | -34.93 | -27.4166
17.5681 | -34.9305 | -27.4241
17.5742 | -34.931 | -27.4319
17.5803 | -34.9314 | -27.4392
17.5864 | -34.9319 | -27.4465
17.5925 | -34.9324 | -27.4544
17.5986 | -34.933 | -27.4618
17.6047 | -34.9335 | -27.4696
17.6108 | -34.934 | -27.477
17.6169 | -34.9344 | -27.4844
17.623 | -34.9349 | -27.492
17.6291 | -34.9354 | -27.4995
17.6352 | -34.9358 | -27.5069
17.6413 | -34.9363 | -27.5146
17.6474 | -34.9368 | -27.5221
17.6535 | -34.9372 | -27.5297
17.6596 | -34.9377 | -27.5372
17.6657 | -34.9382 | -27.5446
17.6718 | -34.9386 | -27.5524
17.6779 | -34.9391 | -27.5599
17.684 | -34.9396 | -27.5676
17.6901 | -34.9401 | -27.5751
17.6963 | -34.9405 | -27.5825
17.7024 | -34.941 | -27.59
17.7085 | -34.9415 | -27.5973
17.7146 | -34.942 | -27.6047
17.7207 | -34.9425 | -27.6121
17.7268 | -34.943 | -27.6196
17.7329 | -34.9434 | -27.6274
17.739 | -34.9439 | -27.6352
17.7451 | -34.9444 | -27.6428
17.7512 | -34.9448 | -27.6504
17.7573 | -34.9453 | -27.6581
17.7634 | -34.9458 | -27.665
17.7695 | -34.9462 | -27.6728
17.7756 | -34.9468 | -27.6802
17.7817 | -34.9472 | -27.6876
17.7878 | -34.9477 | -27.6953
17.7939 | -34.9482 | -27.7028
17.8 | -34.9487 | -27.7102
17.8061 | -34.9491 | -27.7178
17.8122 | -34.9496 | -27.7255
17.8183 | -34.9501 | -27.7331
17.8244 | -34.9506 | -27.7406
17.8305 | -34.9511 | -27.7481
17.8366 | -34.9515 | -27.7553
17.8428 | -34.952 | -27.763
17.8489 | -34.9524 | -27.7703
17.855 | -34.9529 | -27.7776
17.8611 | -34.9534 | -27.785
17.8672 | -34.9539 | -27.7923
17.8733 | -34.9543 | -27.8002
17.8794 | -34.9548 | -27.8079
17.8855 | -34.9553 | -27.8154
17.8916 | -34.9557 | -27.8226
17.8977 | -34.9562 | -27.8301
17.9038 | -34.9567 | -27.8375
17.9099 | -34.9571 | -27.8449
17.916 | -34.9575 | -27.8523
17.9221 | -34.958 | -27.8597
17.9282 | -34.9585 | -27.8672
17.9343 | -34.959 | -27.8749
17.9404 | -34.9594 | -27.8823
17.9465 | -34.9599 | -27.8899
17.9526 | -34.9603 | -27.8978
17.9587 | -34.9608 | -27.9052
17.9648 | -34.9613 | -27.9125
17.9709 | -34.9618 | -27.92
17.977 | -34.9622 | -27.9275
17.9832 | -34.9627 | -27.935
17.9893 | -34.9631 | -27.9425
17.9954 | -34.9636 | -27.9503
18.0015 | -34.9641 | -27.958
18.0076 | -34.9646 | -27.9652
18.0137 | -34.965 | -27.9728
18.0198 | -34.9655 | -27.9802
18.0259 | -34.9659 | -27.9879
18.032 | -34.9663 | -27.9951
18.0381 | -34.9668 | -28.0024
18.0442 | -34.9673 | -28.0101
18.0503 | -34.9678 | -28.0177
18.0564 | -34.9682 | -28.0252
18.0625 | -34.9686 | -28.0328
18.0686 | -34.9691 | -28.0402
18.0747 | -34.9696 | -28.0476
18.0808 | -34.97 | -28.0549
18.0869 | -34.9705 | -28.0623
18.093 | -34.971 | -28.0697
18.0991 | -34.9714 | -28.0773
18.1052 | -34.9719 | -28.085
18.1113 | -34.9724 | -28.0922
18.1174 | -34.9729 | -28.0995
18.1236 | -34.9734 | -28.1071
18.1297 | -34.9739 | -28.1145
18.1358 | -34.9744 | -28.1219
18.1419 | -34.9749 | -28.1295
18.148 | -34.9754 | -28.1369
18.1541 | -34.9759 | -28.1442
18.1602 | -34.9764 | -28.1519
18.1663 | -34.9769 | -28.1591
18.1724 | -34.9774 | -28.1668
18.1785 | -34.9778 | -28.1744
18.1846 | -34.9783 | -28.1817
18.1907 | -34.9787 | -28.1893
18.1968 | -34.9792 | -28.1967
18.2029 | -34.9796 | -28.2037
18.209 | -34.9801 | -28.2115
18.2151 | -34.9806 | -28.2192
18.2212 | -34.9811 | -28.2265
18.2273 | -34.9816 | -28.2339
18.2334 | -34.9821 | -28.2413
18.2395 | -34.9825 | -28.2485
18.2456 | -34.983 | -28.2559
18.2517 | -34.9834 | -28.2633
18.2578 | -34.9838 | -28.2705
18.2639 | -34.9844 | -28.2782
18.2701 | -34.9849 | -28.2856
18.2762 | -34.9854 | -28.2933
18.2823 | -34.9858 | -28.3006
18.2884 | -34.9863 | -28.308
18.2945 | -34.9868 | -28.3153
18.3006 | -34.9873 | -28.3227
18.3067 | -34.9878 | -28.3299
18.3128 | -34.9882 | -28.3372
18.3189 | -34.9887 | -28.3447
18.325 | -34.9892 | -28.3516
18.3311 | -34.9897 | -28.3591
18.3372 | -34.9902 | -28.3662
18.3433 | -34.9906 | -28.3736
18.3494 | -34.991 | -28.3811
18.3555 | -34.9915 | -28.3882
18.3616 | -34.992 | -28.3954
18.3677 | -34.9925 | -28.4024
18.3738 | -34.993 | -28.4098
18.3799 | -34.9935 | -28.4169
18.386 | -34.9939 | -28.4242
18.3921 | -34.9944 | -28.4315
18.3982 | -34.9948 | -28.4387
18.4043 | -34.9953 | -28.4459
18.4105 | -34.9957 | -28.4533
18.4166 | -34.9962 | -28.4604
18.4227 | -34.9966 | -28.4679
18.4288 | -34.9971 | -28.4756
18.4349 | -34.9975 | -28.4831
18.441 | -34.998 | -28.4906
18.4471 | -34.9985 | -28.4979
18.4532 | -34.999 | -28.5053
18.4593 | -34.9995 | -28.5125
18.4654 | -34.9999 | -28.5196
18.4715 | -35.0004 | -28.5268
18.4776 | -35.0009 | -28.5343
18.4837 | -35.0014 | -28.5419
18.4898 | -35.0019 | -28.5491
18.4959 | -35.0024 | -28.5566
18.502 | -35.0029 | -28.5644
18.5081 | -35.0034 | -28.572
18.5142 | -35.0038 | -28.5795
18.5203 | -35.0043 | -28.5867
18.5264 | -35.0047 | -28.5937
18.5325 | -35.0052 | -28.6013
18.5386 | -35.0058 | -28.6089
18.5447 | -35.0062 | -28.6162
18.5508 | -35.0067 | -28.6232
18.557 | -35.0072 | -28.6307
18.5631 | -35.0076 | -28.6381
18.5692 | -35.0081 | -28.6453
18.5753 | -35.0086 | -28.6524
18.5814 | -35.0091 | -28.66
18.5875 | -35.0095 | -28.6677
18.5936 | -35.01 | -28.6754
18.5997 | -35.0105 | -28.6826
18.6058 | -35.011 | -28.6897
18.6119 | -35.0115 | -28.6972
18.618 | -35.012 | -28.7047
18.6241 | -35.0125 | -28.7121
18.6302 | -35.013 | -28.7197
18.6363 | -35.0135 | -28.7268
18.6424 | -35.0139 | -28.7339
18.6485 | -35.0143 | -28.7414
18.6546 | -35.0148 | -28.7484
18.6607 | -35.0153 | -28.7555
18.6668 | -35.0158 | -28.7628
18.6729 | -35.0163 | -28.7701
18.679 | -35.0168 | -28.7771
18.6851 | -35.0172 | -28.7844
18.6912 | -35.0177 | -28.7918
18.6974 | -35.0182 | -28.7989
18.7035 | -35.0187 | -28.806
18.7096 | -35.0192 | -28.8135
18.7157 | -35.0197 | -28.8209
18.7218 | -35.0201 | -28.8282
18.7279 | -35.0206 | -28.8356
18.734 | -35.0211 | -28.843
18.7401 | -35.0216 | -28.8501
18.7462 | -35.0221 | -28.8568
18.7523 | -35.0226 | -28.864
18.7584 | -35.0231 | -28.8715
18.7645 | -35.0236 | -28.8787
18.7706 | -35.0241 | -28.886
18.7767 | -35.0246 | -28.8932
18.7828 | -35.025 | -28.9002
18.7889 | -35.0255 | -28.9072
18.795 | -35.026 | -28.9144
18.8011 | -35.0264 | -28.9216
18.8072 | -35.0269 | -28.9291
18.8133 | -35.0274 | -28.9363
18.8194 | -35.0279 | -28.9434
18.8255 | -35.0284 | -28.9505
18.8316 | -35.0288 | -28.9578
18.8377 | -35.0293 | -28.9649
18.8439 | -35.0298 | -28.9724
18.85 | -35.0302 | -28.9797
18.8561 | -35.0307 | -28.9868
18.8622 | -35.0311 | -28.9942
18.8683 | -35.0316 | -29.0013
18.8744 | -35.0321 | -29.0085
18.8805 | -35.0326 | -29.0159
18.8866 | -35.0331 | -29.0234
18.8927 | -35.0336 | -29.0303
18.8988 | -35.0341 | -29.0373
18.9049 | -35.0346 | -29.0447
18.911 | -35.0351 | -29.0516
18.9171 | -35.0356 | -29.0589
18.9232 | -35.0361 | -29.0659
18.9293 | -35.0365 | -29.0733
18.9354 | -35.037 | -29.0804
18.9415 | -35.0375 | -29.0872
18.9476 | -35.038 | -29.0945
18.9537 | -35.0384 | -29.1021
18.9598 | -35.039 | -29.1086
18.9659 | -35.0395 | -29.1156
18.972 | -35.04 | -29.1231
18.9781 | -35.0405 | -29.1303
18.9843 | -35.041 | -29.1374
18.9904 | -35.0414 | -29.1446
18.9965 | -35.0419 | -29.1515
19.0026 | -35.0424 | -29.1586
19.0087 | -35.0429 | -29.1656
19.0148 | -35.0434 | -29.1724
19.0209 | -35.0439 | -29.1798
19.027 | -35.0444 | -29.1872
19.0331 | -35.0448 | -29.1944
19.0392 | -35.0453 | -29.2016
19.0453 | -35.0458 | -29.2089
19.0514 | -35.0463 | -29.2158
19.0575 | -35.0467 | -29.223
19.0636 | -35.0472 | -29.2296
19.0697 | -35.0478 | -29.2372
19.0758 | -35.0482 | -29.244
19.0819 | -35.0487 | -29.2515
19.088 | -35.0492 | -29.2583
19.0941 | -35.0497 | -29.2654
19.1002 | -35.0502 | -29.2724
19.1063 | -35.0507 | -29.2793
19.1124 | -35.0513 | -29.2864
19.1185 | -35.0517 | -29.2937
19.1246 | -35.0521 | -29.3007
19.1308 | -35.0526 | -29.3074
19.1369 | -35.0531 | -29.3143
19.143 | -35.0535 | -29.3217
19.1491 | -35.054 | -29.3289
19.1552 | -35.0544 | -29.3364
19.1613 | -35.055 | -29.3436
19.1674 | -35.0554 | -29.3509
19.1735 | -35.0559 | -29.3578
19.1796 | -35.0564 | -29.3648
19.1857 | -35.0569 | -29.3722
19.1918 | -35.0573 | -29.3793
19.1979 | -35.0578 | -29.3865
19.204 | -35.0583 | -29.3935
19.2101 | -35.0588 | -29.4007
19.2162 | -35.0592 | -29.4079
19.2223 | -35.0597 | -29.4145
19.2284 | -35.0603 | -29.422
19.2345 | -35.0607 | -29.4293
19.2406 | -35.0613 | -29.4364
19.2467 | -35.0617 | -29.4436
19.2528 | -35.0622 | -29.4508
19.2589 | -35.0626 | -29.4579
19.265 | -35.0632 | -29.4648
19.2712 | -35.0636 | -29.4719
19.2773 | -35.0641 | -29.4787
19.2834 | -35.0646 | -29.4855
19.2895 | -35.0651 | -29.4929
19.2956 | -35.0656 | -29.5002
19.3017 | -35.0661 | -29.5072
19.3078 | -35.0666 | -29.5145
19.3139 | -35.067 | -29.5216
19.32 | -35.0676 | -29.5284
19.3261 | -35.068 | -29.5354
19.3322 | -35.0685 | -29.5426
19.3383 | -35.069 | -29.5495
19.3444 | -35.0695 | -29.5561
19.3505 | -35.0699 | -29.5628
19.3566 | -35.0704 | -29.5697
19.3627 | -35.0709 | -29.5769
19.3688 | -35.0714 | -29.5839
19.3749 | -35.0718 | -29.5911
19.381 | -35.0724 | -29.5976
19.3871 | -35.0729 | -29.6047
19.3932 | -35.0733 | -29.6123
19.3993 | -35.0738 | -29.6195
19.4054 | -35.0743 | -29.6268
19.4115 | -35.0748 | -29.634
19.4177 | -35.0752 | -29.6407
19.4238 | -35.0757 | -29.6475
19.4299 | -35.0763 | -29.6544
19.436 | -35.0767 | -29.6619
19.4421 | -35.0773 | -29.6685
19.4482 | -35.0778 | -29.6757
19.4543 | -35.0783 | -29.6823
19.4604 | -35.0788 | -29.6892
19.4665 | -35.0793 | -29.6962
19.4726 | -35.0798 | -29.7031
19.4787 | -35.0802 | -29.71
19.4848 | -35.0807 | -29.7169
19.4909 | -35.0813 | -29.7237
19.497 | -35.0817 | -29.7304
19.5031 | -35.0822 | -29.7371
19.5092 | -35.0826 | -29.7441
19.5153 | -35.0831 | -29.751
19.5214 | -35.0837 | -29.7574
19.5275 | -35.0842 | -29.7641
19.5336 | -35.0846 | -29.7711
19.5397 | -35.0851 | -29.7779
19.5458 | -35.0856 | -29.7849
19.5519 | -35.0861 | -29.7918
19.5581 | -35.0866 | -29.7986
19.5642 | -35.0871 | -29.8053
19.5703 | -35.0876 | -29.8123
19.5764 | -35.0881 | -29.8194
19.5825 | -35.0886 | -29.8264
19.5886 | -35.089 | -29.8331
19.5947 | -35.0896 | -29.84
19.6008 | -35.09 | -29.8469
19.6069 | -35.0905 | -29.8536
19.613 | -35.0909 | -29.8605
19.6191 | -35.0914 | -29.8676
19.6252 | -35.0919 | -29.8747
19.6313 | -35.0924 | -29.881
19.6374 | -35.0928 | -29.8883
19.6435 | -35.0933 | -29.8953
19.6496 | -35.0938 | -29.9021
19.6557 | -35.0943 | -29.909
19.6618 | -35.0948 | -29.9158
19.6679 | -35.0953 | -29.9226
19.674 | -35.0958 | -29.9295
19.6801 | -35.0962 | -29.9364
19.6862 | -35.0968 | -29.9433
19.6923 | -35.0973 | -29.9498
19.6984 | -35.0978 | -29.9562
19.7046 | -35.0983 | -29.9633
19.7107 | -35.0988 | -29.97
19.7168 | -35.0993 | -29.9769
19.7229 | -35.0997 | -29.984
19.729 | -35.1002 | -29.9909
19.7351 | -35.1006 | -29.9974
19.7412 | -35.1011 | -30.0044
19.7473 | -35.1017 | -30.0114
19.7534 | -35.1021 | -30.0185
19.7595 | -35.1027 | -30.0255
19.7656 | -35.1032 | -30.0327
19.7717 | -35.1037 | -30.0391
19.7778 | -35.1041 | -30.0463
19.7839 | -35.1046 | -30.0529
19.79 | -35.1051 | -30.0599
19.7961 | -35.1056 | -30.0667
19.8022 | -35.106 | -30.0732
19.8083 | -35.1066 | -30.0801
19.8144 | -35.107 | -30.0869
19.8205 | -35.1075 | -30.0938
19.8266 | -35.1081 | -30.1007
19.8327 | -35.1085 | -30.1075
19.8388 | -35.109 | -30.1146
19.845 | -35.1095 | -30.1212
19.8511 | -35.11 | -30.1283
19.8572 | -35.1105 | -30.1352
19.8633 | -35.111 | -30.142
19.8694 | -35.1115 | -30.149
19.8755 | -35.112 | -30.1556
19.8816 | -35.1126 | -30.1625
19.8877 | -35.1131 | -30.1693
19.8938 | -35.1135 | -30.1759
19.8999 | -35.114 | -30.1827
19.906 | -35.1145 | -30.1897
19.9121 | -35.115 | -30.196
19.9182 | -35.1155 | -30.2027
19.9243 | -35.116 | -30.2098
19.9304 | -35.1164 | -30.2168
19.9365 | -35.1169 | -30.2237
19.9426 | -35.1174 | -30.2303
19.9487 | -35.1179 | -30.2369
19.9548 | -35.1184 | -30.2438
19.9609 | -35.1189 | -30.2506
19.967 | -35.1195 | -30.2571
19.9731 | -35.1199 | -30.2639
19.9792 | -35.1204 | -30.2702
19.9853 | -35.1209 | -30.2772
19.9915 | -35.1214 | -30.2841
19.9976 | -35.1219 | -30.2909
20.0037 | -35.1223 | -30.2978
20.0098 | -35.1228 | -30.3048
20.0159 | -35.1233 | -30.3111
20.022 | -35.1238 | -30.3183
20.0281 | -35.1243 | -30.3251
20.0342 | -35.1248 | -30.3319
20.0403 | -35.1253 | -30.3384
20.0464 | -35.1258 | -30.3455
20.0525 | -35.1263 | -30.3523
20.0586 | -35.1268 | -30.3589
20.0647 | -35.1274 | -30.3661
20.0708 | -35.1279 | -30.3727
20.0769 | -35.1283 | -30.3793
20.083 | -35.1289 | -30.3857
20.0891 | -35.1293 | -30.3925
20.0952 | -35.1299 | -30.3992
20.1013 | -35.1303 | -30.406
20.1074 | -35.1308 | -30.4127
20.1135 | -35.1313 | -30.4192
20.1196 | -35.1318 | -30.4257
20.1257 | -35.1323 | -30.4328
20.1319 | -35.1328 | -30.4396
20.138 | -35.1333 | -30.4461
20.1441 | -35.1338 | -30.4531
20.1502 | -35.1343 | -30.4601
20.1563 | -35.1348 | -30.4663
20.1624 | -35.1353 | -30.4731
20.1685 | -35.1359 | -30.4799
20.1746 | -35.1364 | -30.4864
20.1807 | -35.1368 | -30.4931
20.1868 | -35.1374 | -30.5
20.1929 | -35.1379 | -30.5063
20.199 | -35.1384 | -30.5133
20.2051 | -35.1389 | -30.5196
20.2112 | -35.1394 | -30.526
20.2173 | -35.1398 | -30.5329
20.2234 | -35.1403 | -30.54
20.2295 | -35.1408 | -30.5467
20.2356 | -35.1413 | -30.5536
20.2417 | -35.1418 | -30.5601
20.2478 | -35.1423 | -30.5669
20.2539 | -35.1428 | -30.574
20.26 | -35.1433 | -30.5803
20.2661 | -35.1438 | -30.5873
20.2723 | -35.1443 | -30.594
20.2784 | -35.1447 | -30.6008
20.2845 | -35.1452 | -30.6073
20.2906 | -35.1457 | -30.6139
20.2967 | -35.1462 | -30.6201
20.3028 | -35.1467 | -30.6268
20.3089 | -35.1471 | -30.6338
20.315 | -35.1476 | -30.6405
20.3211 | -35.1481 | -30.6473
20.3272 | -35.1487 | -30.6537
20.3333 | -35.1491 | -30.6602
20.3394 | -35.1496 | -30.6669
20.3455 | -35.1501 | -30.6735
20.3516 | -35.1506 | -30.6801
20.3577 | -35.1511 | -30.6865
20.3638 | -35.1516 | -30.6932
20.3699 | -35.1521 | -30.6995
20.376 | -35.1527 | -30.7062
20.3821 | -35.1532 | -30.7128
20.3882 | -35.1537 | -30.7195
20.3943 | -35.1542 | -30.7262
20.4004 | -35.1547 | -30.7327
20.4065 | -35.1552 | -30.7388
20.4126 | -35.1557 | -30.7454
20.4188 | -35.1562 | -30.7522
20.4249 | -35.1567 | -30.7589
20.431 | -35.1572 | -30.7655
20.4371 | -35.1577 | -30.7714
20.4432 | -35.1581 | -30.7782
20.4493 | -35.1586 | -30.7847
20.4554 | -35.1591 | -30.7913
20.4615 | -35.1596 | -30.7982
20.4676 | -35.16 | -30.8049
20.4737 | -35.1605 | -30.8112
20.4798 | -35.161 | -30.8179
20.4859 | -35.1615 | -30.8244
20.492 | -35.162 | -30.8311
20.4981 | -35.1625 | -30.8375
20.5042 | -35.163 | -30.8438
20.5103 | -35.1635 | -30.8503
20.5164 | -35.164 | -30.8568
20.5225 | -35.1645 | -30.8633
20.5286 | -35.165 | -30.8701
20.5347 | -35.1655 | -30.8761
20.5408 | -35.1659 | -30.8828
20.5469 | -35.1665 | -30.8893
20.553 | -35.167 | -30.8962
20.5592 | -35.1675 | -30.9025
20.5653 | -35.1678 | -30.9093
20.5714 | -35.1683 | -30.9152
20.5775 | -35.1688 | -30.9215
20.5836 | -35.1692 | -30.9279
20.5897 | -35.1698 | -30.9345
20.5958 | -35.1703 | -30.9414
20.6019 | -35.1707 | -30.9479
20.608 | -35.1712 | -30.9546
20.6141 | -35.1717 | -30.9607
20.6202 | -35.1721 | -30.9674
20.6263 | -35.1726 | -30.974
20.6324 | -35.1731 | -30.9805
20.6385 | -35.1735 | -30.9868
20.6446 | -35.174 | -30.9937
20.6507 | -35.1745 | -30.9999
20.6568 | -35.1749 | -31.0063
20.6629 | -35.1754 | -31.0128
20.669 | -35.1759 | -31.0194
20.6751 | -35.1763 | -31.026
20.6812 | -35.1767 | -31.0319
20.6873 | -35.1773 | -31.0383
20.6934 | -35.1778 | -31.045
20.6995 | -35.1782 | -31.0518
20.7057 | -35.1787 | -31.0583
20.7118 | -35.1792 | -31.0644
20.7179 | -35.1797 | -31.0709
20.724 | -35.1802 | -31.0774
20.7301 | -35.1806 | -31.0839
20.7362 | -35.1811 | -31.0905
20.7423 | -35.1817 | -31.0968
20.7484 | -35.1821 | -31.1038
20.7545 | -35.1826 | -31.1102
20.7606 | -35.1832 | -31.1172
20.7667 | -35.1836 | -31.1242
20.7728 | -35.1842 | -31.1303
20.7789 | -35.1847 | -31.1368
20.785 | -35.1851 | -31.1436
20.7911 | -35.1855 | -31.1498
20.7972 | -35.186 | -31.1565
20.8033 | -35.1864 | -31.1627
20.8094 | -35.1869 | -31.1698
20.8155 | -35.1873 | -31.1758
20.8216 | -35.1877 | -31.1826
20.8277 | -35.1882 | -31.1887
20.8338 | -35.1886 | -31.195
20.8399 | -35.1891 | -31.2017
20.8461 | -35.1897 | -31.2081
20.8522 | -35.1902 | -31.2146
20.8583 | -35.1908 | -31.2204
20.8644 | -35.1913 | -31.2266
20.8705 | -35.1917 | -31.2334
20.8766 | -35.1921 | -31.2401
20.8827 | -35.1926 | -31.2462
20.8888 | -35.1931 | -31.2524
20.8949 | -35.1936 | -31.2588
20.901 | -35.1941 | -31.2651
20.9071 | -35.1947 | -31.2715
20.9132 | -35.1951 | -31.2774
20.9193 | -35.1956 | -31.2836
20.9254 | -35.1961 | -31.2899
20.9315 | -35.1966 | -31.2964
20.9376 | -35.197 | -31.3028
20.9437 | -35.1976 | -31.3094
20.9498 | -35.1981 | -31.316
20.9559 | -35.1986 | -31.3223
20.962 | -35.1991 | -31.3292
20.9681 | -35.1996 | -31.3352
20.9742 | -35.2001 | -31.3418
20.9803 | -35.2006 | -31.3487
20.9864 | -35.201 | -31.3553
20.9926 | -35.2014 | -31.3617
20.9987 | -35.2019 | -31.3684
21.0048 | -35.2024 | -31.3746
21.0109 | -35.2028 | -31.3809
21.017 | -35.2033 | -31.3875
21.0231 | -35.2038 | -31.394
21.0292 | -35.2043 | -31.4001
21.0353 | -35.2048 | -31.4063
21.0414 | -35.2052 | -31.4127
21.0475 | -35.2057 | -31.419
21.0536 | -35.2062 | -31.4254
21.0597 | -35.2067 | -31.4315
21.0658 | -35.2071 | -31.4376
21.0719 | -35.2077 | -31.4432
21.078 | -35.2082 | -31.45
21.0841 | -35.2086 | -31.4573
21.0902 | -35.2091 | -31.4644
21.0963 | -35.2096 | -31.4715
21.1024 | -35.2101 | -31.4778
21.1085 | -35.2105 | -31.4851
21.1146 | -35.211 | -31.4919
21.1207 | -35.2114 | -31.4986
21.1268 | -35.2119 | -31.5047
21.133 | -35.2124 | -31.5111
21.1391 | -35.213 | -31.5174
21.1452 | -35.2134 | -31.524
21.1513 | -35.2139 | -31.5304
21.1574 | -35.2144 | -31.5368
21.1635 | -35.2149 | -31.5431
21.1696 | -35.2154 | -31.5494
21.1757 | -35.2158 | -31.5555
21.1818 | -35.2163 | -31.562
21.1879 | -35.2169 | -31.5679
21.194 | -35.2174 | -31.5742
21.2001 | -35.218 | -31.5801
21.2062 | -35.2184 | -31.5859
21.2123 | -35.2188 | -31.5916
21.2184 | -35.2193 | -31.5982
21.2245 | -35.2197 | -31.6046
21.2306 | -35.2201 | -31.6112
21.2367 | -35.2206 | -31.6179
21.2428 | -35.2211 | -31.6237
21.2489 | -35.2216 | -31.6298
21.255 | -35.222 | -31.6366
21.2611 | -35.2225 | -31.6428
21.2672 | -35.223 | -31.6493
21.2733 | -35.2234 | -31.6562
21.2795 | -35.2239 | -31.6623
21.2856 | -35.2243 | -31.6688
21.2917 | -35.2248 | -31.6753
21.2978 | -35.2252 | -31.682
21.3039 | -35.2257 | -31.6882
21.31 | -35.2261 | -31.6945
21.3161 | -35.2266 | -31.6996
21.3222 | -35.227 | -31.7058
21.3283 | -35.2274 | -31.7122
21.3344 | -35.2279 | -31.7183
21.3405 | -35.2282 | -31.7248
21.3466 | -35.2287 | -31.7315
21.3527 | -35.2291 | -31.7381
21.3588 | -35.2296 | -31.7444
21.3649 | -35.2301 | -31.7508
21.371 | -35.2306 | -31.7569
21.3771 | -35.231 | -31.7632
21.3832 | -35.2315 | -31.769
21.3893 | -35.2319 | -31.775
21.3954 | -35.2323 | -31.7814
21.4015 | -35.2328 | -31.7877
21.4076 | -35.2333 | -31.7938
21.4137 | -35.2336 | -31.8005
21.4199 | -35.2341 | -31.8076
21.426 | -35.2345 | -31.814
21.4321 | -35.2351 | -31.8201
21.4382 | -35.2355 | -31.8264
21.4443 | -35.236 | -31.8328
21.4504 | -35.2365 | -31.8392
21.4565 | -35.237 | -31.8452
21.4626 | -35.2375 | -31.8513
21.4687 | -35.238 | -31.8575
21.4748 | -35.2385 | -31.864
21.4809 | -35.239 | -31.8703
21.487 | -35.2395 | -31.8765
21.4931 | -35.2399 | -31.8824
21.4992 | -35.2405 | -31.8885
21.5053 | -35.2409 | -31.8955
21.5114 | -35.2414 | -31.9025
21.5175 | -35.2419 | -31.909
21.5236 | -35.2422 | -31.9149
21.5297 | -35.2429 | -31.9214
21.5358 | -35.2433 | -31.9277
21.5419 | -35.2437 | -31.9344
21.548 | -35.2442 | -31.9404
21.5541 | -35.2445 | -31.9466
21.5602 | -35.2449 | -31.953
21.5664 | -35.2453 | -31.9597
21.5725 | -35.2457 | -31.9659
21.5786 | -35.2462 | -31.9724
21.5847 | -35.2466 | -31.9782
21.5908 | -35.247 | -31.9844
21.5969 | -35.2475 | -31.9898
21.603 | -35.248 | -31.9962
21.6091 | -35.2486 | -32.0029
21.6152 | -35.249 | -32.0095
21.6213 | -35.2494 | -32.0158
21.6274 | -35.2499 | -32.0223
21.6335 | -35.2503 | -32.0288
21.6396 | -35.2508 | -32.036
21.6457 | -35.2514 | -32.043
21.6518 | -35.2517 | -32.0491
21.6579 | -35.2522 | -32.0549
21.664 | -35.2527 | -32.0616
21.6701 | -35.2532 | -32.0679
21.6762 | -35.2536 | -32.0744
21.6823 | -35.254 | -32.0811
21.6884 | -35.2545 | -32.0875
21.6945 | -35.255 | -32.094
21.7006 | -35.2554 | -32.1
21.7068 | -35.2558 | -32.1057
21.7129 | -35.2563 | -32.1119
21.719 | -35.2568 | -32.1181
21.7251 | -35.2572 | -32.1244
21.7312 | -35.2577 | -32.1308
21.7373 | -35.2581 | -32.1375
21.7434 | -35.2586 | -32.1439
21.7495 | -35.259 | -32.1502
21.7556 | -35.2595 | -32.1566
21.7617 | -35.26 | -32.163
21.7678 | -35.2605 | -32.1697
21.7739 | -35.2609 | -32.1757
21.78 | -35.2614 | -32.182
21.7861 | -35.2618 | -32.1885
21.7922 | -35.2623 | -32.195
21.7983 | -35.2628 | -32.2019
21.8044 | -35.2633 | -32.208
21.8105 | -35.2637 | -32.2135
21.8166 | -35.2642 | -32.2202
21.8227 | -35.2646 | -32.2266
21.8288 | -35.2651 | -32.2328
21.8349 | -35.2655 | -32.2395
21.841 | -35.266 | -32.2459
21.8471 | -35.2664 | -32.2528
21.8533 | -35.2668 | -32.2588
21.8594 | -35.2672 | -32.2652
21.8655 | -35.2677 | -32.271
21.8716 | -35.2682 | -32.2774
21.8777 | -35.2686 | -32.2836
21.8838 | -35.2691 | -32.2898
21.8899 | -35.2695 | -32.2962
21.896 | -35.2701 | -32.3025
21.9021 | -35.2705 | -32.3089
21.9082 | -35.2709 | -32.3153
21.9143 | -35.2713 | -32.3215
21.9204 | -35.2718 | -32.3277
21.9265 | -35.2721 | -32.3336
21.9326 | -35.2727 | -32.34
21.9387 | -35.2732 | -32.3462
21.9448 | -35.2736 | -32.3522
21.9509 | -35.2741 | -32.3585
21.957 | -35.2746 | -32.3646
21.9631 | -35.275 | -32.3711
21.9692 | -35.2755 | -32.3778
21.9753 | -35.2759 | -32.3837
21.9814 | -35.2763 | -32.39
21.9875 | -35.2768 | -32.3964
21.9937 | -35.2772 | -32.4029
21.9998 | -35.2777 | -32.4093
22.0059 | -35.2781 | -32.4153
22.012 | -35.2785 | -32.421
22.0181 | -35.279 | -32.4272
22.0242 | -35.2794 | -32.4336
22.0303 | -35.2799 | -32.4401
22.0364 | -35.2803 | -32.4465
22.0425 | -35.2808 | -32.4524
22.0486 | -35.2812 | -32.4586
22.0547 | -35.2816 | -32.4648
22.0608 | -35.282 | -32.4711
22.0669 | -35.2824 | -32.4775
22.073 | -35.2828 | -32.4831
22.0791 | -35.2832 | -32.4896
22.0852 | -35.2837 | -32.4954
22.0913 | -35.2842 | -32.5018
22.0974 | -35.2846 | -32.508
22.1035 | -35.285 | -32.5139
22.1096 | -35.2855 | -32.5207
22.1157 | -35.2859 | -32.527
22.1218 | -35.2864 | -32.534
22.1279 | -35.2867 | -32.5395
22.134 | -35.2872 | -32.546
22.1402 | -35.2877 | -32.5518
22.1463 | -35.2882 | -32.5586
22.1524 | -35.2887 | -32.5652
22.1585 | -35.2891 | -32.5711
22.1646 | -35.2895 | -32.577
22.1707 | -35.29 | -32.5838
22.1768 | -35.2904 | -32.5899
22.1829 | -35.2909 | -32.5966
22.189 | -35.2913 | -32.6031
22.1951 | -35.2919 | -32.6094
22.2012 | -35.2923 | -32.6156
22.2073 | -35.2928 | -32.6219
22.2134 | -35.2932 | -32.6284
22.2195 | -35.2936 | -32.6345
22.2256 | -35.2941 | -32.6407
22.2317 | -35.2945 | -32.6466
22.2378 | -35.2949 | -32.6525
22.2439 | -35.2954 | -32.6585
22.25 | -35.2958 | -32.6646
22.2561 | -35.2962 | -32.6709
22.2622 | -35.2966 | -32.6771
22.2683 | -35.297 | -32.6829
22.2744 | -35.2975 | -32.6891
22.2806 | -35.2979 | -32.6952
22.2867 | -35.2983 | -32.7015
22.2928 | -35.2988 | -32.7077
22.2989 | -35.2992 | -32.7141
22.305 | -35.2995 | -32.7203
22.3111 | -35.3 | -32.7265
22.3172 | -35.3004 | -32.733
22.3233 | -35.3009 | -32.739
22.3294 | -35.3013 | -32.7456
22.3355 | -35.3018 | -32.7513
22.3416 | -35.3023 | -32.7574
22.3477 | -35.3027 | -32.7637
22.3538 | -35.3031 | -32.7702
22.3599 | -35.3036 | -32.7759
22.366 | -35.304 | -32.7821
22.3721 | -35.3044 | -32.788
22.3782 | -35.3049 | -32.7947
22.3843 | -35.3053 | -32.8007
22.3904 | -35.3058 | -32.8074
22.3965 | -35.3062 | -32.8142
22.4026 | -35.3067 | -32.8203
22.4087 | -35.3069 | -32.8263
22.4148 | -35.3074 | -32.8331
22.4209 | -35.3078 | -32.8394
22.4271 | -35.3083 | -32.8458
22.4332 | -35.3087 | -32.8514
22.4393 | -35.3091 | -32.8575
22.4454 | -35.3095 | -32.8637
22.4515 | -35.3098 | -32.8701
22.4576 | -35.3103 | -32.8763
22.4637 | -35.3107 | -32.8828
22.4698 | -35.3112 | -32.8891
22.4759 | -35.3115 | -32.8951
22.482 | -35.312 | -32.9017
22.4881 | -35.3123 | -32.9077
22.4942 | -35.3128 | -32.914
22.5003 | -35.3131 | -32.9204
22.5064 | -35.3136 | -32.9265
22.5125 | -35.314 | -32.9327
22.5186 | -35.3145 | -32.9391
22.5247 | -35.3149 | -32.9454
22.5308 | -35.3153 | -32.9521
22.5369 | -35.3157 | -32.9584
22.543 | -35.3161 | -32.9643
22.5491 | -35.3165 | -32.9704
22.5552 | -35.317 | -32.9764
22.5613 | -35.3174 | -32.9829
22.5675 | -35.3178 | -32.9892
22.5736 | -35.3182 | -32.9953
22.5797 | -35.3186 | -33.0019
22.5858 | -35.3189 | -33.0083
22.5919 | -35.3193 | -33.0145
22.598 | -35.3197 | -33.0207
22.6041 | -35.3201 | -33.027
22.6102 | -35.3206 | -33.0334
22.6163 | -35.321 | -33.0395
22.6224 | -35.3215 | -33.0458
22.6285 | -35.3219 | -33.0523
22.6346 | -35.3223 | -33.0577
22.6407 | -35.3227 | -33.0643
22.6468 | -35.3231 | -33.0708
22.6529 | -35.3235 | -33.0768
22.659 | -35.3239 | -33.0832
22.6651 | -35.3243 | -33.0893
22.6712 | -35.3247 | -33.0958
22.6773 | -35.3251 | -33.1017
22.6834 | -35.3254 | -33.1079
22.6895 | -35.3258 | -33.1144
22.6956 | -35.3263 | -33.1203
22.7017 | -35.3266 | -33.1265
22.7079 | -35.3271 | -33.1327
22.714 | -35.3275 | -33.1387
22.7201 | -35.3279 | -33.1447
22.7262 | -35.3283 | -33.1512
22.7323 | -35.3287 | -33.1571
22.7384 | -35.329 | -33.1634
22.7445 | -35.3295 | -33.1697
22.7506 | -35.3299 | -33.1757
22.7567 | -35.3303 | -33.1818
22.7628 | -35.3307 | -33.1884
22.7689 | -35.3311 | -33.1943
22.775 | -35.3315 | -33.2006
22.7811 | -35.332 | -33.2064
22.7872 | -35.3324 | -33.2128
22.7933 | -35.3329 | -33.219
22.7994 | -35.3334 | -33.2254
22.8055 | -35.3338 | -33.2317
22.8116 | -35.3342 | -33.238
22.8177 | -35.3346 | -33.2441
22.8238 | -35.3349 | -33.2504
22.8299 | -35.3352 | -33.2563
22.836 | -35.3355 | -33.2633
22.8421 | -35.3359 | -33.269
22.8482 | -35.3363 | -33.2751
22.8544 | -35.3367 | -33.2814
22.8605 | -35.3372 | -33.2876
22.8666 | -35.3375 | -33.2943
22.8727 | -35.3378 | -33.3004
22.8788 | -35.3382 | -33.3064
22.8849 | -35.3387 | -33.3124
22.891 | -35.3391 | -33.3188
22.8971 | -35.3396 | -33.3248
22.9032 | -35.3399 | -33.3314
22.9093 | -35.3403 | -33.337
22.9154 | -35.3407 | -33.3431
22.9215 | -35.3411 | -33.3489
22.9276 | -35.3415 | -33.3551
22.9337 | -35.3419 | -33.3612
22.9398 | -35.3422 | -33.3667
22.9459 | -35.3426 | -33.3733
22.952 | -35.343 | -33.3798
22.9581 | -35.3434 | -33.3862
22.9642 | -35.3438 | -33.3924
22.9703 | -35.3441 | -33.3984
22.9764 | -35.3445 | -33.4044
22.9825 | -35.3448 | -33.4105
22.9886 | -35.3453 | -33.4171
22.9948 | -35.3457 | -33.4237
23.0009 | -35.346 | -33.4299
23.007 | -35.3463 | -33.436
23.0131 | -35.3467 | -33.442
23.0192 | -35.3471 | -33.4482
23.0253 | -35.3475 | -33.4548
23.0314 | -35.348 | -33.4604
23.0375 | -35.3484 | -33.467
23.0436 | -35.3487 | -33.4727
23.0497 | -35.3491 | -33.4786
23.0558 | -35.3496 | -33.4844
23.0619 | -35.3499 | -33.4904
23.068 | -35.3503 | -33.4968
23.0741 | -35.3508 | -33.5027
23.0802 | -35.3511 | -33.5088
23.0863 | -35.3515 | -33.5152
23.0924 | -35.3519 | -33.5209
23.0985 | -35.3522 | -33.5268
23.1046 | -35.3527 | -33.5325
23.1107 | -35.3531 | -33.539
23.1168 | -35.3535 | -33.5447
23.1229 | -35.354 | -33.5508
23.129 | -35.3543 | -33.5573
23.1351 | -35.3547 | -33.5631
23.1413 | -35.3551 | -33.569
23.1474 | -35.3554 | -33.5753
23.1535 | -35.3558 | -33.5812
23.1596 | -35.3561 | -33.5868
23.1657 | -35.3565 | -33.5932
23.1718 | -35.357 | -33.5986
23.1779 | -35.3574 | -33.6047
23.184 | -35.3575 | -33.6112
23.1901 | -35.358 | -33.6178
23.1962 | -35.3584 | -33.6239
23.2023 | -35.3588 | -33.6295
23.2084 | -35.3592 | -33.6357
23.2145 | -35.3596 | -33.6417
23.2206 | -35.36 | -33.6478
23.2267 | -35.3604 | -33.6543
23.2328 | -35.3607 | -33.6602
23.2389 | -35.3611 | -33.6665
23.245 | -35.3615 | -33.6727
23.2511 | -35.3619 | -33.6787
23.2572 | -35.3623 | -33.6847
23.2633 | -35.3628 | -33.6911
23.2694 | -35.3632 | -33.6976
23.2755 | -35.3636 | -33.7036
23.2817 | -35.364 | -33.7097
23.2878 | -35.3644 | -33.7155
23.2939 | -35.3647 | -33.7214
23.3 | -35.3651 | -33.7279
23.3061 | -35.3655 | -33.7337
23.3122 | -35.3659 | -33.7402
23.3183 | -35.3663 | -33.7461
23.3244 | -35.3666 | -33.7523
23.3305 | -35.367 | -33.7584
23.3366 | -35.3673 | -33.7645
23.3427 | -35.3677 | -33.7701
23.3488 | -35.3681 | -33.7768
23.3549 | -35.3684 | -33.7828
23.361 | -35.3688 | -33.7891
23.3671 | -35.3692 | -33.7951
23.3732 | -35.3696 | -33.8014
23.3793 | -35.37 | -33.8074
23.3854 | -35.3704 | -33.8134
23.3915 | -35.3707 | -33.8199
23.3976 | -35.371 | -33.8264
23.4037 | -35.3714 | -33.8322
23.4098 | -35.3718 | -33.8388
23.4159 | -35.3722 | -33.845
23.422 | -35.3725 | -33.8509
23.4282 | -35.3728 | -33.8567
23.4343 | -35.3732 | -33.8627
23.4404 | -35.3736 | -33.8697
23.4465 | -35.374 | -33.8756
23.4526 | -35.3745 | -33.8819
23.4587 | -35.3748 | -33.8881
23.4648 | -35.375 | -33.8942
23.4709 | -35.3753 | -33.8998
23.477 | -35.3757 | -33.9061
23.4831 | -35.3761 | -33.9117
23.4892 | -35.3765 | -33.918
23.4953 | -35.377 | -33.9244
23.5014 | -35.3772 | -33.9301
23.5075 | -35.3777 | -33.9362
23.5136 | -35.3781 | -33.9423
23.5197 | -35.3783 | -33.9484
23.5258 | -35.3787 | -33.9547
23.5319 | -35.379 | -33.9609
23.538 | -35.3795 | -33.9672
23.5441 | -35.3798 | -33.9734
23.5502 | -35.3803 | -33.9794
23.5563 | -35.3806 | -33.9856
23.5624 | -35.381 | -33.9915
23.5686 | -35.3813 | -33.9968
23.5747 | -35.3816 | -34.0032
23.5808 | -35.382 | -34.009
23.5869 | -35.3823 | -34.0146
23.593 | -35.3827 | -34.021
23.5991 | -35.383 | -34.027
23.6052 | -35.3833 | -34.0338
23.6113 | -35.3837 | -34.04
23.6174 | -35.384 | -34.0464
23.6235 | -35.3844 | -34.0522
23.6296 | -35.3847 | -34.0587
23.6357 | -35.385 | -34.0656
23.6418 | -35.3855 | -34.0713
23.6479 | -35.3858 | -34.0773
23.654 | -35.3861 | -34.0831
23.6601 | -35.3865 | -34.089
23.6662 | -35.3868 | -34.0952
23.6723 | -35.3871 | -34.101
23.6784 | -35.3875 | -34.1073
23.6845 | -35.3878 | -34.113
23.6906 | -35.3882 | -34.1191
23.6967 | -35.3885 | -34.125
23.7028 | -35.3889 | -34.1312
23.7089 | -35.3892 | -34.1378
23.7151 | -35.3896 | -34.1435
23.7212 | -35.39 | -34.1494
23.7273 | -35.3903 | -34.1554
23.7334 | -35.3907 | -34.1618
23.7395 | -35.3911 | -34.1677
23.7456 | -35.3914 | -34.1733
23.7517 | -35.3917 | -34.1794
23.7578 | -35.3922 | -34.1856
23.7639 | -35.3925 | -34.192
23.77 | -35.3928 | -34.1979
23.7761 | -35.3932 | -34.2037
23.7822 | -35.3935 | -34.2103
23.7883 | -35.3939 | -34.2165
23.7944 | -35.3942 | -34.2222
23.8005 | -35.3946 | -34.2286
23.8066 | -35.3951 | -34.2343
23.8127 | -35.3955 | -34.2409
23.8188 | -35.3959 | -34.2471
23.8249 | -35.3963 | -34.2537
23.831 | -35.3966 | -34.2594
23.8371 | -35.3968 | -34.2652
23.8432 | -35.3972 | -34.2711
23.8493 | -35.3975 | -34.2766
23.8555 | -35.3979 | -34.282
23.8616 | -35.3983 | -34.2885
23.8677 | -35.3987 | -34.2946
23.8738 | -35.399 | -34.301
23.8799 | -35.3993 | -34.3071
23.886 | -35.3996 | -34.3134
23.8921 | -35.4 | -34.3199
23.8982 | -35.4004 | -34.3264
23.9043 | -35.4008 | -34.3326
23.9104 | -35.4011 | -34.339
23.9165 | -35.4015 | -34.345
23.9226 | -35.4019 | -34.3504
23.9287 | -35.4023 | -34.3567
23.9348 | -35.4025 | -34.3628
23.9409 | -35.4028 | -34.3687
23.947 | -35.4032 | -34.3747
23.9531 | -35.4035 | -34.3806
23.9592 | -35.4039 | -34.3864
23.9653 | -35.4043 | -34.3928
23.9714 | -35.4046 | -34.3987
23.9775 | -35.4049 | -34.4048
23.9836 | -35.4052 | -34.4108
23.9897 | -35.4056 | -34.4164
23.9958 | -35.406 | -34.4226
24.002 | -35.4063 | -34.4289
24.0081 | -35.4066 | -34.4348
24.0142 | -35.407 | -34.4411
24.0203 | -35.4073 | -34.447
24.0264 | -35.4076 | -34.4527
24.0325 | -35.4079 | -34.4592
24.0386 | -35.4082 | -34.4649
24.0447 | -35.4086 | -34.4709
24.0508 | -35.409 | -34.4768
24.0569 | -35.4093 | -34.4828
24.063 | -35.4095 | -34.4895
24.0691 | -35.4099 | -34.495
24.0752 | -35.4103 | -34.5009
24.0813 | -35.4107 | -34.5073
24.0874 | -35.411 | -34.5126
24.0935 | -35.4113 | -34.5188
24.0996 | -35.4116 | -34.5256
24.1057 | -35.412 | -34.5312
24.1118 | -35.4123 | -34.537
24.1179 | -35.4126 | -34.5423
24.124 | -35.4131 | -34.5486
24.1301 | -35.4134 | -34.5542
24.1362 | -35.4138 | -34.5599
24.1424 | -35.4142 | -34.5655
24.1485 | -35.4145 | -34.5712
24.1546 | -35.4148 | -34.5777
24.1607 | -35.4151 | -34.5843
24.1668 | -35.4155 | -34.5897
24.1729 | -35.4157 | -34.5955
24.179 | -35.416 | -34.6014
24.1851 | -35.4164 | -34.6074
24.1912 | -35.4168 | -34.6136
24.1973 | -35.4172 | -34.6197
24.2034 | -35.4174 | -34.6252
24.2095 | -35.4178 | -34.6313
24.2156 | -35.4182 | -34.6373
24.2217 | -35.4186 | -34.6428
24.2278 | -35.419 | -34.6485
24.2339 | -35.4194 | -34.6542
24.24 | -35.4196 | -34.6599
24.2461 | -35.4199 | -34.6663
24.2522 | -35.4203 | -34.672
24.2583 | -35.4206 | -34.6784
24.2644 | -35.4209 | -34.6839
24.2705 | -35.4212 | -34.6903
24.2766 | -35.4215 | -34.6962
24.2827 | -35.4218 | -34.7023
24.2889 | -35.4221 | -34.7078
24.295 | -35.4225 | -34.7138
24.3011 | -35.4229 | -34.72
24.3072 | -35.4232 | -34.7261
24.3133 | -35.4236 | -34.7323
24.3194 | -35.4239 | -34.7381
24.3255 | -35.4242 | -34.744
24.3316 | -35.4245 | -34.7502
24.3377 | -35.4248 | -34.7562
24.3438 | -35.4252 | -34.7617
24.3499 | -35.4255 | -34.7682
24.356 | -35.4258 | -34.7742
24.3621 | -35.4261 | -34.7809
24.3682 | -35.4264 | -34.787
24.3743 | -35.4267 | -34.7927
24.3804 | -35.427 | -34.7983
24.3865 | -35.4273 | -34.8039
24.3926 | -35.4277 | -34.8103
24.3987 | -35.428 | -34.8162
24.4048 | -35.4283 | -34.8217
24.4109 | -35.4287 | -34.8277
24.417 | -35.4289 | -34.8335
24.4231 | -35.4293 | -34.84
24.4293 | -35.4296 | -34.8449
24.4354 | -35.43 | -34.8509
24.4415 | -35.4302 | -34.857
24.4476 | -35.4306 | -34.8623
24.4537 | -35.4309 | -34.8681
24.4598 | -35.4313 | -34.8739
24.4659 | -35.4317 | -34.8801
24.472 | -35.4319 | -34.8863
24.4781 | -35.4323 | -34.8922
24.4842 | -35.4325 | -34.8976
24.4903 | -35.4328 | -34.9038
24.4964 | -35.4331 | -34.9093
24.5025 | -35.4335 | -34.9154
24.5086 | -35.4338 | -34.9215
24.5147 | -35.4342 | -34.927
24.5208 | -35.4346 | -34.9328
24.5269 | -35.4349 | -34.9389
24.533 | -35.4353 | -34.945
24.5391 | -35.4356 | -34.9512
24.5452 | -35.4359 | -34.9571
24.5513 | -35.4362 | -34.963
24.5574 | -35.4366 | -34.9693
24.5635 | -35.4369 | -34.9755
24.5696 | -35.4373 | -34.9812
24.5758 | -35.4376 | -34.9871
24.5819 | -35.4379 | -34.9932
24.588 | -35.4384 | -34.9992
24.5941 | -35.4387 | -35.0049
24.6002 | -35.4391 | -35.0109
24.6063 | -35.4394 | -35.0166
24.6124 | -35.4397 | -35.0226
24.6185 | -35.4399 | -35.0285
24.6246 | -35.4403 | -35.0341
24.6307 | -35.4406 | -35.0403
24.6368 | -35.441 | -35.0468
24.6429 | -35.4413 | -35.0527
24.649 | -35.4416 | -35.0584
24.6551 | -35.4419 | -35.0639
24.6612 | -35.4422 | -35.0696
24.6673 | -35.4425 | -35.0763
24.6734 | -35.4428 | -35.0821
24.6795 | -35.4432 | -35.088
24.6856 | -35.4435 | -35.0936
24.6917 | -35.4439 | -35.0994
24.6978 | -35.4442 | -35.1048
24.7039 | -35.4444 | -35.111
24.71 | -35.4448 | -35.1169
24.7162 | -35.4452 | -35.1227
24.7223 | -35.4455 | -35.1286
24.7284 | -35.446 | -35.1337
24.7345 | -35.4462 | -35.1394
24.7406 | -35.4466 | -35.1455
24.7467 | -35.4469 | -35.1515
24.7528 | -35.4471 | -35.1568
24.7589 | -35.4475 | -35.1626
24.765 | -35.4479 | -35.1681
24.7711 | -35.4482 | -35.1741
24.7772 | -35.4486 | -35.1802
24.7833 | -35.449 | -35.186
24.7894 | -35.4493 | -35.1917
24.7955 | -35.4495 | -35.1977
24.8016 | -35.4497 | -35.2036
24.8077 | -35.45 | -35.2093
24.8138 | -35.4504 | -35.215
24.8199 | -35.4507 | -35.221
24.826 | -35.451 | -35.2264
24.8321 | -35.4514 | -35.2322
24.8382 | -35.4518 | -35.2371
24.8443 | -35.4521 | -35.2433
24.8504 | -35.4525 | -35.249
24.8565 | -35.4528 | -35.2549
24.8627 | -35.4531 | -35.261
24.8688 | -35.4534 | -35.2662
24.8749 | -35.4537 | -35.2717
24.881 | -35.454 | -35.2774
24.8871 | -35.4543 | -35.2828
24.8932 | -35.4546 | -35.2888
24.8993 | -35.4549 | -35.2945
24.9054 | -35.4553 | -35.2999
24.9115 | -35.4556 | -35.3062
24.9176 | -35.4559 | -35.3118
24.9237 | -35.4563 | -35.3176
24.9298 | -35.4566 | -35.3228
24.9359 | -35.4569 | -35.3287
24.942 | -35.4572 | -35.3341
24.9481 | -35.4575 | -35.34
24.9542 | -35.4579 | -35.3463
24.9603 | -35.4582 | -35.3522
24.9664 | -35.4584 | -35.3581
24.9725 | -35.4588 | -35.364
24.9786 | -35.4591 | -35.3698
24.9847 | -35.4595 | -35.3756
24.9908 | -35.4598 | -35.381
24.9969 | -35.4601 | -35.3872
25.0031 | -35.4604 | -35.393
25.0092 | -35.4607 | -35.3992
25.0153 | -35.4611 | -35.4048
25.0214 | -35.4613 | -35.4107
25.0275 | -35.4616 | -35.4166
25.0336 | -35.462 | -35.4226
25.0397 | -35.4623 | -35.4286
25.0458 | -35.4626 | -35.4342
25.0519 | -35.4629 | -35.4399
25.058 | -35.4632 | -35.4456
25.0641 | -35.4635 | -35.4514
25.0702 | -35.4638 | -35.457
25.0763 | -35.4642 | -35.4628
25.0824 | -35.4645 | -35.4688
25.0885 | -35.4648 | -35.4745
25.0946 | -35.4651 | -35.4798
25.1007 | -35.4655 | -35.4852
25.1068 | -35.4659 | -35.4906
25.1129 | -35.4662 | -35.4962
25.119 | -35.4665 | -35.5018
25.1251 | -35.4668 | -35.5076
25.1312 | -35.4672 | -35.5133
25.1373 | -35.4675 | -35.519
25.1435 | -35.4678 | -35.5254
25.1496 | -35.4681 | -35.5312
25.1557 | -35.4684 | -35.537
25.1618 | -35.4689 | -35.543
25.1679 | -35.4692 | -35.5485
25.174 | -35.4696 | -35.5542
25.1801 | -35.4698 | -35.5599
25.1862 | -35.4701 | -35.5654
25.1923 | -35.4704 | -35.571
25.1984 | -35.4707 | -35.5765
25.2045 | -35.4711 | -35.5817
25.2106 | -35.4714 | -35.587
25.2167 | -35.4716 | -35.5929
25.2228 | -35.4718 | -35.5978
25.2289 | -35.4722 | -35.6031
25.235 | -35.4725 | -35.6091
25.2411 | -35.4728 | -35.6148
25.2472 | -35.4731 | -35.6206
25.2533 | -35.4733 | -35.626
25.2594 | -35.4737 | -35.6314
25.2655 | -35.4741 | -35.6371
25.2716 | -35.4744 | -35.6427
25.2777 | -35.4747 | -35.6482
25.2838 | -35.4751 | -35.6541
25.29 | -35.4753 | -35.6598
25.2961 | -35.4757 | -35.6651
25.3022 | -35.476 | -35.6706
25.3083 | -35.4763 | -35.676
25.3144 | -35.4766 | -35.6818
25.3205 | -35.4769 | -35.6875
25.3266 | -35.4771 | -35.6934
25.3327 | -35.4774 | -35.6999
25.3388 | -35.4776 | -35.7056
25.3449 | -35.478 | -35.7111
25.351 | -35.4782 | -35.717
25.3571 | -35.4786 | -35.7227
25.3632 | -35.4789 | -35.7285
25.3693 | -35.4793 | -35.7343
25.3754 | -35.4797 | -35.7397
25.3815 | -35.4801 | -35.7457
25.3876 | -35.4804 | -35.7517
25.3937 | -35.4807 | -35.757
25.3998 | -35.481 | -35.7622
25.4059 | -35.4814 | -35.768
25.412 | -35.4817 | -35.7731
25.4181 | -35.4821 | -35.779
25.4242 | -35.4823 | -35.7851
25.4304 | -35.4826 | -35.7907
25.4365 | -35.483 | -35.7965
25.4426 | -35.4833 | -35.802
25.4487 | -35.4837 | -35.8079
25.4548 | -35.4839 | -35.814
25.4609 | -35.4842 | -35.8196
25.467 | -35.4847 | -35.8247
25.4731 | -35.485 | -35.8307
25.4792 | -35.4853 | -35.8367
25.4853 | -35.4857 | -35.8422
25.4914 | -35.486 | -35.8478
25.4975 | -35.4863 | -35.8535
25.5036 | -35.4868 | -35.8592
25.5097 | -35.4872 | -35.8642
25.5158 | -35.4875 | -35.8698
25.5219 | -35.4878 | -35.8753
25.528 | -35.4882 | -35.8807
25.5341 | -35.4885 | -35.8864
25.5402 | -35.4888 | -35.8928
25.5463 | -35.4891 | -35.8983
25.5524 | -35.4894 | -35.904
25.5585 | -35.4897 | -35.9094
25.5646 | -35.49 | -35.9154
25.5707 | -35.4903 | -35.9211
25.5769 | -35.4907 | -35.9268
25.583 | -35.491 | -35.9329
25.5891 | -35.4912 | -35.9385
25.5952 | -35.4916 | -35.9443
25.6013 | -35.4918 | -35.9501
25.6074 | -35.4921 | -35.9549
25.6135 | -35.4924 | -35.9608
25.6196 | -35.4927 | -35.9669
25.6257 | -35.493 | -35.9728
25.6318 | -35.4933 | -35.978
25.6379 | -35.4936 | -35.9836
25.644 | -35.4939 | -35.9891
25.6501 | -35.4942 | -35.9943
25.6562 | -35.4945 | -35.9999
25.6623 | -35.4949 | -36.0054
25.6684 | -35.4952 | -36.0107
25.6745 | -35.4955 | -36.0167
25.6806 | -35.4959 | -36.0217
25.6867 | -35.4962 | -36.0277
25.6928 | -35.4966 | -36.0335
25.6989 | -35.4968 | -36.0393
25.705 | -35.4971 | -36.0451
25.7111 | -35.4974 | -36.0507
25.7173 | -35.4976 | -36.0568
25.7234 | -35.498 | -36.0624
25.7295 | -35.4983 | -36.0682
25.7356 | -35.4986 | -36.0742
25.7417 | -35.4989 | -36.0797
25.7478 | -35.4993 | -36.0851
25.7539 | -35.4996 | -36.0909
25.76 | -35.5 | -36.0966
25.7661 | -35.5002 | -36.1023
25.7722 | -35.5006 | -36.1081
25.7783 | -35.5009 | -36.1138
25.7844 | -35.5012 | -36.1196
25.7905 | -35.5015 | -36.1252
25.7966 | -35.5018 | -36.1307
25.8027 | -35.5021 | -36.1366
25.8088 | -35.5024 | -36.1427
25.8149 | -35.5028 | -36.1483
25.821 | -35.503 | -36.1533
25.8271 | -35.5034 | -36.1589
25.8332 | -35.5037 | -36.1644
25.8393 | -35.504 | -36.1703
25.8454 | -35.5043 | -36.176
25.8515 | -35.5047 | -36.1822
25.8576 | -35.505 | -36.188
25.8638 | -35.5054 | -36.1935
25.8699 | -35.5057 | -36.1992
25.876 | -35.506 | -36.2053
25.8821 | -35.5063 | -36.2111
25.8882 | -35.5066 | -36.2164
25.8943 | -35.507 | -36.2218
25.9004 | -35.5072 | -36.2279
25.9065 | -35.5075 | -36.2332
25.9126 | -35.5078 | -36.239
25.9187 | -35.5082 | -36.2451
25.9248 | -35.5085 | -36.2502
25.9309 | -35.5088 | -36.2559
25.937 | -35.5091 | -36.2615
25.9431 | -35.5094 | -36.2673
25.9492 | -35.5097 | -36.2728
25.9553 | -35.51 | -36.2783
25.9614 | -35.5103 | -36.2839
25.9675 | -35.5106 | -36.2892
25.9736 | -35.5109 | -36.2951
25.9797 | -35.5112 | -36.301
25.9858 | -35.5116 | -36.3067
25.9919 | -35.5119 | -36.3123
25.998 | -35.5122 | -36.3175
26.0042 | -35.5125 | -36.323
26.0103 | -35.5128 | -36.3287
26.0164 | -35.5131 | -36.3342
26.0225 | -35.5135 | -36.3399
26.0286 | -35.5138 | -36.346
26.0347 | -35.5142 | -36.3514
26.0408 | -35.5145 | -36.3568
26.0469 | -35.5148 | -36.3622
26.053 | -35.515 | -36.368
26.0591 | -35.5153 | -36.3739
26.0652 | -35.5157 | -36.3794
26.0713 | -35.516 | -36.3847
26.0774 | -35.5163 | -36.3902
26.0835 | -35.5167 | -36.3953
26.0896 | -35.517 | -36.4007
26.0957 | -35.5174 | -36.4063
26.1018 | -35.5177 | -36.412
26.1079 | -35.518 | -36.4171
26.114 | -35.5183 | -36.4227
26.1201 | -35.5186 | -36.428
26.1262 | -35.5189 | -36.4336
26.1323 | -35.5193 | -36.4391
26.1384 | -35.5196 | -36.4444
26.1445 | -35.5199 | -36.45
26.1507 | -35.5202 | -36.4555
26.1568 | -35.5205 | -36.4609
26.1629 | -35.5208 | -36.4665
26.169 | -35.5211 | -36.472
26.1751 | -35.5215 | -36.4771
26.1812 | -35.5218 | -36.4825
26.1873 | -35.5222 | -36.4881
26.1934 | -35.5224 | -36.4933
26.1995 | -35.5228 | -36.4985
26.2056 | -35.5231 | -36.504
26.2117 | -35.5234 | -36.5095
26.2178 | -35.5237 | -36.5151
26.2239 | -35.5241 | -36.5206
26.23 | -35.5244 | -36.5267
26.2361 | -35.5247 | -36.5322
26.2422 | -35.525 | -36.5374
26.2483 | -35.5254 | -36.5431
26.2544 | -35.5257 | -36.5487
26.2605 | -35.526 | -36.5544
26.2666 | -35.5263 | -36.5602
26.2727 | -35.5266 | -36.5661
26.2788 | -35.527 | -36.5715
26.2849 | -35.5273 | -36.5771
26.2911 | -35.5276 | -36.5827
26.2972 | -35.5279 | -36.5882
26.3033 | -35.5283 | -36.594
26.3094 | -35.5286 | -36.5992
26.3155 | -35.5289 | -36.6045
26.3216 | -35.5292 | -36.6098
26.3277 | -35.5294 | -36.6152
26.3338 | -35.5297 | -36.6207
26.3399 | -35.53 | -36.6259
26.346 | -35.5303 | -36.6314
26.3521 | -35.5307 | -36.6368
26.3582 | -35.531 | -36.6426
26.3643 | -35.5314 | -36.648
26.3704 | -35.5317 | -36.6535
26.3765 | -35.5321 | -36.6588
26.3826 | -35.5324 | -36.6645
26.3887 | -35.5327 | -36.6701
26.3948 | -35.533 | -36.6756
26.4009 | -35.5332 | -36.6811
26.407 | -35.5336 | -36.6865
26.4131 | -35.5339 | -36.6919
26.4192 | -35.5342 | -36.6975
26.4253 | -35.5345 | -36.7032
26.4314 | -35.5348 | -36.7084
26.4376 | -35.5351 | -36.714
26.4437 | -35.5354 | -36.7191
26.4498 | -35.5358 | -36.7248
26.4559 | -35.5361 | -36.7295
26.462 | -35.5365 | -36.7352
26.4681 | -35.5368 | -36.7405
26.4742 | -35.5371 | -36.7462
26.4803 | -35.5374 | -36.7516
26.4864 | -35.5377 | -36.7571
26.4925 | -35.538 | -36.7626
26.4986 | -35.5384 | -36.7681
26.5047 | -35.5387 | -36.7739
26.5108 | -35.539 | -36.7792
26.5169 | -35.5393 | -36.7846
26.523 | -35.5396 | -36.7901
26.5291 | -35.5398 | -36.7952
26.5352 | -35.5401 | -36.8008
26.5413 | -35.5404 | -36.8066
26.5474 | -35.5407 | -36.8124
26.5535 | -35.5411 | -36.8181
26.5596 | -35.5414 | -36.8231
26.5657 | -35.5417 | -36.8287
26.5718 | -35.542 | -36.8343
26.578 | -35.5424 | -36.8398
26.5841 | -35.5427 | -36.8455
26.5902 | -35.543 | -36.8507
26.5963 | -35.5433 | -36.856
26.6024 | -35.5437 | -36.8615
26.6085 | -35.544 | -36.8671
26.6146 | -35.5443 | -36.8726
26.6207 | -35.5446 | -36.8783
26.6268 | -35.5449 | -36.8838
26.6329 | -35.5453 | -36.8892
26.639 | -35.5456 | -36.895
26.6451 | -35.5458 | -36.9004
26.6512 | -35.5461 | -36.9058
26.6573 | -35.5464 | -36.9115
26.6634 | -35.5468 | -36.917
26.6695 | -35.5471 | -36.9225
26.6756 | -35.5474 | -36.9282
26.6817 | -35.5478 | -36.9338
26.6878 | -35.5481 | -36.9395
26.6939 | -35.5484 | -36.9454
26.7 | -35.5487 | -36.9507
26.7061 | -35.5491 | -36.9564
26.7122 | -35.5494 | -36.9622
26.7183 | -35.5497 | -36.968
26.7245 | -35.55 | -36.9736
26.7306 | -35.5503 | -36.9795
26.7367 | -35.5507 | -36.985
26.7428 | -35.551 | -36.9907
26.7489 | -35.5513 | -36.9967
26.755 | -35.5517 | -37.0021
26.7611 | -35.5521 | -37.0075
26.7672 | -35.5524 | -37.0128
26.7733 | -35.5527 | -37.0181
26.7794 | -35.5531 | -37.0234
26.7855 | -35.5534 | -37.0288
26.7916 | -35.5538 | -37.0342
26.7977 | -35.5541 | -37.04
26.8038 | -35.5544 | -37.0454
26.8099 | -35.5548 | -37.051
26.816 | -35.5551 | -37.0567
26.8221 | -35.5554 | -37.0617
26.8282 | -35.5558 | -37.0671
26.8343 | -35.5561 | -37.0728
26.8404 | -35.5564 | -37.0783
26.8465 | -35.5567 | -37.0837
26.8526 | -35.557 | -37.0896
26.8587 | -35.5574 | -37.095
26.8649 | -35.5577 | -37.1009
26.871 | -35.558 | -37.1064
26.8771 | -35.5583 | -37.112
26.8832 | -35.5586 | -37.1174
26.8893 | -35.559 | -37.1227
26.8954 | -35.5593 | -37.1278
26.9015 | -35.5596 | -37.1333
26.9076 | -35.5599 | -37.139
26.9137 | -35.5603 | -37.1445
26.9198 | -35.5607 | -37.1499
26.9259 | -35.561 | -37.1551
26.932 | -35.5613 | -37.1605
26.9381 | -35.5616 | -37.1661
26.9442 | -35.5618 | -37.1715
26.9503 | -35.5622 | -37.1771
26.9564 | -35.5625 | -37.1823
26.9625 | -35.5628 | -37.1877
26.9686 | -35.5632 | -37.1933
26.9747 | -35.5636 | -37.1992
26.9808 | -35.5639 | -37.2047
26.9869 | -35.5641 | -37.21
26.993 | -35.5645 | -37.2158
26.9991 | -35.5648 | -37.2213
27.0052 | -35.5651 | -37.2269
27.0114 | -35.5655 | -37.2326
27.0175 | -35.5658 | -37.2384
27.0236 | -35.5661 | -37.244
27.0297 | -35.5664 | -37.2497
27.0358 | -35.5668 | -37.2551
27.0419 | -35.5671 | -37.2608
27.048 | -35.5674 | -37.2661
27.0541 | -35.5677 | -37.2716
27.0602 | -35.568 | -37.2771
27.0663 | -35.5684 | -37.283
27.0724 | -35.5687 | -37.2886
27.0785 | -35.5691 | -37.2938
27.0846 | -35.5694 | -37.2992
27.0907 | -35.5697 | -37.305
27.0968 | -35.57 | -37.3103
27.1029 | -35.5703 | -37.316
27.109 | -35.5707 | -37.3212
27.1151 | -35.571 | -37.3265
27.1212 | -35.5713 | -37.3313
27.1273 | -35.5716 | -37.3367
27.1334 | -35.5719 | -37.3423
27.1395 | -35.5723 | -37.3474
27.1456 | -35.5726 | -37.3531
27.1518 | -35.5729 | -37.3586
27.1579 | -35.5732 | -37.3639
27.164 | -35.5736 | -37.3694
27.1701 | -35.5739 | -37.3748
27.1762 | -35.5742 | -37.3805
27.1823 | -35.5746 | -37.3861
27.1884 | -35.5749 | -37.3915
27.1945 | -35.5753 | -37.3969
27.2006 | -35.5756 | -37.4022
27.2067 | -35.5759 | -37.4075
27.2128 | -35.5762 | -37.4128
27.2189 | -35.5765 | -37.4182
27.225 | -35.5768 | -37.4234
27.2311 | -35.5771 | -37.4287
27.2372 | -35.5775 | -37.4342
27.2433 | -35.5778 | -37.4396
27.2494 | -35.5781 | -37.4449
27.2555 | -35.5784 | -37.4501
27.2616 | -35.5788 | -37.4557
27.2677 | -35.5792 | -37.4613
27.2738 | -35.5794 | -37.4667
27.2799 | -35.5798 | -37.4723
27.286 | -35.5801 | -37.4777
27.2921 | -35.5804 | -37.4833
27.2983 | -35.5807 | -37.4886
27.3044 | -35.5811 | -37.4942
27.3105 | -35.5814 | -37.4998
27.3166 | -35.5817 | -37.5049
27.3227 | -35.5821 | -37.5104
27.3288 | -35.5824 | -37.5158
27.3349 | -35.5828 | -37.5209
27.341 | -35.5831 | -37.5264
27.3471 | -35.5834 | -37.5316
27.3532 | -35.5837 | -37.5371
27.3593 | -35.5841 | -37.5426
27.3654 | -35.5844 | -37.5483
27.3715 | -35.5847 | -37.5536
27.3776 | -35.5851 | -37.5592
27.3837 | -35.5854 | -37.5646
27.3898 | -35.5857 | -37.5699
27.3959 | -35.5861 | -37.5753
27.402 | -35.5864 | -37.5808
27.4081 | -35.5868 | -37.5861
27.4142 | -35.5871 | -37.5915
27.4203 | -35.5874 | -37.5965
27.4264 | -35.5877 | -37.6015
27.4325 | -35.588 | -37.6069
27.4387 | -35.5883 | -37.6124
27.4448 | -35.5887 | -37.618
27.4509 | -35.589 | -37.6238
27.457 | -35.5894 | -37.6292
27.4631 | -35.5897 | -37.6348
27.4692 | -35.5899 | -37.6399
27.4753 | -35.5902 | -37.6452
27.4814 | -35.5905 | -37.6507
27.4875 | -35.5908 | -37.6564
27.4936 | -35.5912 | -37.6621
27.4997 | -35.5915 | -37.6676
27.5058 | -35.5918 | -37.6728
27.5119 | -35.5922 | -37.6783
27.518 | -35.5925 | -37.6838
27.5241 | -35.5928 | -37.6891
27.5302 | -35.5931 | -37.6944
27.5363 | -35.5935 | -37.7002
27.5424 | -35.5938 | -37.7054
27.5485 | -35.5941 | -37.7106
27.5546 | -35.5944 | -37.7157
27.5607 | -35.5948 | -37.7206
27.5668 | -35.595 | -37.7258
27.5729 | -35.5953 | -37.7312
27.5791 | -35.5957 | -37.7365
27.5852 | -35.596 | -37.7417
27.5913 | -35.5964 | -37.7473
27.5974 | -35.5966 | -37.753
27.6035 | -35.5969 | -37.7584
27.6096 | -35.5973 | -37.7635
27.6157 | -35.5976 | -37.7689
27.6218 | -35.5979 | -37.7743
27.6279 | -35.5982 | -37.7794
27.634 | -35.5985 | -37.7851
27.6401 | -35.5988 | -37.7905
27.6462 | -35.5992 | -37.7957
27.6523 | -35.5995 | -37.8006
27.6584 | -35.5998 | -37.8061
27.6645 | -35.6002 | -37.8113
27.6706 | -35.6006 | -37.8166
27.6767 | -35.6009 | -37.8219
27.6828 | -35.6012 | -37.8269
27.6889 | -35.6015 | -37.8325
27.695 | -35.6018 | -37.8378
27.7011 | -35.6021 | -37.8432
27.7072 | -35.6024 | -37.849
27.7133 | -35.6028 | -37.8541
27.7194 | -35.6031 | -37.8596
27.7256 | -35.6034 | -37.865
27.7317 | -35.6037 | -37.8705
27.7378 | -35.604 | -37.8759
27.7439 | -35.6044 | -37.8813
27.75 | -35.6047 | -37.8866
27.7561 | -35.6051 | -37.8916
27.7622 | -35.6054 | -37.8971
27.7683 | -35.6057 | -37.9028
27.7744 | -35.606 | -37.9083
27.7805 | -35.6063 | -37.9138
27.7866 | -35.6066 | -37.9193
27.7927 | -35.6069 | -37.9243
27.7988 | -35.6072 | -37.9297
27.8049 | -35.6074 | -37.9348
27.811 | -35.6078 | -37.9399
27.8171 | -35.6081 | -37.9452
27.8232 | -35.6085 | -37.9504
27.8293 | -35.6088 | -37.9559
27.8354 | -35.6091 | -37.9614
27.8415 | -35.6095 | -37.9666
27.8476 | -35.6098 | -37.972
27.8537 | -35.6102 | -37.9774
27.8598 | -35.6105 | -37.9828
27.866 | -35.6108 | -37.9881
27.8721 | -35.6112 | -37.9934
27.8782 | -35.6115 | -37.9985
27.8843 | -35.6118 | -38.0038
27.8904 | -35.6121 | -38.0091
27.8965 | -35.6125 | -38.014
27.9026 | -35.6128 | -38.0191
27.9087 | -35.6131 | -38.0245
27.9148 | -35.6134 | -38.0298
27.9209 | -35.6137 | -38.035
27.927 | -35.6141 | -38.04
27.9331 | -35.6145 | -38.0457
27.9392 | -35.6148 | -38.0513
27.9453 | -35.615 | -38.0565
27.9514 | -35.6154 | -38.0615
27.9575 | -35.6157 | -38.0667
27.9636 | -35.616 | -38.0721
27.9697 | -35.6163 | -38.0776
27.9758 | -35.6166 | -38.0831
27.9819 | -35.617 | -38.0888
27.988 | -35.6173 | -38.0939
27.9941 | -35.6176 | -38.0995
28.0002 | -35.6179 | -38.1045
28.0063 | -35.6182 | -38.1098
28.0125 | -35.6185 | -38.1153
28.0186 | -35.6188 | -38.1207
28.0247 | -35.6192 | -38.1256
28.0308 | -35.6195 | -38.1308
28.0369 | -35.6198 | -38.1356
28.043 | -35.6202 | -38.1414
28.0491 | -35.6205 | -38.1467
28.0552 | -35.6207 | -38.1519
28.0613 | -35.6211 | -38.1571
28.0674 | -35.6214 | -38.1622
28.0735 | -35.6217 | -38.1677

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4259 | -42.2541
21.4908 | -35.5144 | -35.1313
16.4948 | -34.9025 | -27.7253
12.6602 | -34.5457 | -21.4401
9.7171 | -34.3321 | -16.5255
7.4582 | -34.1818 | -12.5446
5.7244 | -34.0772 | -9.3199
4.3936 | -34.0124 | -6.7184
3.3722 | -34 | -4.666
2.5883 | -33.9859 | -3.1138
1.9866 | -33.9832 | -1.8733
1.5248 | -33.9815 | -0.87922
1.1703 | -33.9978 | -0.23456
1 | -33.9845 | 0.086902
0.89822 | -33.979 | 0.35629
0.68941 | -33.9585 | 0.9342
0.52919 | -33.9703 | 1.3783
0.40612 | -33.9725 | 1.9837
0.31171 | -33.972 | 2.4281
0.23928 | -33.9622 | 3.1564
0.18358 | -33.9717 | 3.5604
0.14091 | -33.9873 | 4.4321
0.10819 | -34.0061 | 5.5866
0.083017 | -34.0307 | 7.0635
0.063705 | -34.0665 | 8.9402
0.048923 | -34.1521 | 11.437
0.037479 | -34.2506 | 14.6267
0.028801 | -34.434 | 18.6406
0.022078 | -34.7348 | 23.9516
0.017023 | -35.2141 | 30.361
0.013018 | -36.0164 | 38.0527
0.010014 | -37.1341 | 45.9453

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003689 | -6.7545
0.007377 | -24.6578
0.011066 | -34.2495
0.014754 | -41.0842
0.018443 | -45.5171
0.022131 | -47.3737
0.02582 | -48.4275
0.029508 | -49.6233
0.033197 | -51.0331
0.036885 | -51.4501
0.040574 | -53.0515
0.044262 | -55.3334
0.047951 | -55.807
0.051639 | -56.2705
0.055328 | -56.9758
0.059016 | -58.5571
0.062705 | -59.5388
0.066393 | -58.9468
0.070082 | -59.3864
0.07377 | -60.4279
0.077459 | -61.2827
0.081147 | -59.512
0.084836 | -60.4793
0.088524 | -62.8564
0.092213 | -62.8148
0.095901 | -62.6083
0.09959 | -62.8462
0.10328 | -63.0464
0.10697 | -64.5959
0.11066 | -66.3188
0.11434 | -64.8744
0.11803 | -63.684
0.12172 | -63.8007
0.12541 | -64.2208
0.1291 | -65.6584
0.13279 | -66.7162
0.13648 | -66.9248
0.14016 | -68.0457
0.14385 | -67.5243
0.14754 | -67.4677
0.15123 | -67.4155
0.15492 | -67.0562
0.15861 | -66.8139
0.16229 | -68.7529
0.16598 | -69.1138
0.16967 | -68.9567
0.17336 | -69.8825
0.17705 | -70.9091
0.18074 | -71.4524
0.18443 | -71.0676
0.18811 | -69.304
0.1918 | -69.5174
0.19549 | -69.789
0.19918 | -69.4729
0.20287 | -70.0855
0.20656 | -71.302
0.21024 | -72.0038
0.21393 | -71.7211
0.21762 | -72.0727
0.22131 | -71.6424
0.225 | -70.3677
0.22869 | -71.6285
0.23238 | -72.8788
0.23606 | -73.1647
0.23975 | -72.7129
0.24344 | -72.8255
0.24713 | -73.1306
0.25082 | -73.2452
0.25451 | -73.7257
0.2582 | -74.5023
0.26188 | -73.8704
0.26557 | -73.5213
0.26926 | -74.2194
0.27295 | -74.735
0.27664 | -75.3543
0.28033 | -74.7086
0.28402 | -74.5746
0.2877 | -74.152
0.29139 | -73.382
0.29508 | -74.2889
0.29877 | -74.8739
0.30246 | -74.9632
0.30615 | -75.1745
0.30983 | -75.2977
0.31352 | -75.9824
0.31721 | -76.4103
0.3209 | -75.4848
0.32459 | -75.1492
0.32828 | -75.847
0.33197 | -75.7263
0.33565 | -75.4366
0.33934 | -75.2964
0.34303 | -75.4833
0.34672 | -77.1051
0.35041 | -77.4296
0.3541 | -77.5795
0.35779 | -77.2216
0.36147 | -77.2931
0.36516 | -77.5852
0.36885 | -76.5808
0.37254 | -76.6957
0.37623 | -77.2317
0.37992 | -76.3291
0.3836 | -75.9554
0.38729 | -76.4613
0.39098 | -77.8316
0.39467 | -77.5231
0.39836 | -78.1712
0.40205 | -78.4622
0.40574 | -78.2451
0.40942 | -78.2905
0.41311 | -77.4731
0.4168 | -77.8833
0.42049 | -78.4324
0.42418 | -80.1544
0.42787 | -79.3773
0.43156 | -78.5988
0.43524 | -78.9459
0.43893 | -78.5014
0.44262 | -78.2206
0.44631 | -78.761
0.45 | -78.6133
0.45369 | -78.4629
0.45737 | -78.8113
0.46106 | -79.6704
0.46475 | -79.2642
0.46844 | -79.3562
0.47213 | -79.4635
0.47582 | -79.4542
0.47951 | -79.3427
0.48319 | -79.528
0.48688 | -80.1813
0.49057 | -80.5098
0.49426 | -80.0605
0.49795 | -80.9993
0.50164 | -79.9367
0.50533 | -80.744
0.50901 | -80.6097
0.5127 | -80.1382
0.51639 | -81.0218
0.52008 | -81.5477
0.52377 | -80.5797
0.52746 | -79.8604
0.53114 | -79.5792
0.53483 | -79.6373
0.53852 | -79.7212
0.54221 | -79.9131
0.5459 | -80.1029
0.54959 | -80.4452
0.55328 | -80.9757
0.55696 | -80.9654
0.56065 | -80.6503
0.56434 | -80.7329
0.56803 | -80.7153
0.57172 | -80.5275
0.57541 | -80.9088
0.5791 | -81.6132
0.58278 | -81.4901
0.58647 | -81.1187
0.59016 | -81.1186
0.59385 | -81.357
0.59754 | -81.3439
0.60123 | -80.7545
0.60491 | -80.1911
0.6086 | -80.4
0.61229 | -81.2164
0.61598 | -82.1695
0.61967 | -82.6522
0.62336 | -82.4988
0.62705 | -81.9809
0.63073 | -81.5178
0.63442 | -81.3806
0.63811 | -81.9922
0.6418 | -82.62
0.64549 | -82.3378
0.64918 | -82.0072
0.65287 | -82.0904
0.65655 | -82.2032
0.66024 | -82.2066
0.66393 | -82.2801
0.66762 | -82.3907
0.67131 | -82.2841
0.675 | -81.8815
0.67868 | -81.8807
0.68237 | -82.4419
0.68606 | -83.1901
0.68975 | -83.1545
0.69344 | -82.9089
0.69713 | -83.174
0.70082 | -83.8128
0.70451 | -84.3509
0.70819 | -84.2324
0.71188 | -83.4903
0.71557 | -82.9786
0.71926 | -83.0907
0.72295 | -83.3368
0.72664 | -83.3157
0.73032 | -83.345
0.73401 | -83.5721
0.7377 | -83.5731
0.74139 | -83.5401
0.74508 | -83.7478
0.74877 | -83.6177
0.75245 | -83.4216
0.75614 | -83.4772
0.75983 | -83.8317
0.76352 | -84.2107
0.76721 | -83.9169
0.7709 | -83.0758
0.77459 | -83.0142
0.77828 | -83.6808
0.78196 | -84.1683
0.78565 | -84.2173
0.78934 | -84.213
0.79303 | -84.286
0.79672 | -84.4672
0.80041 | -84.7129
0.80409 | -84.933
0.80778 | -85.2304
0.81147 | -85.255
0.81516 | -85.0082
0.81885 | -85.0456
0.82254 | -84.9808
0.82622 | -84.4649
0.82991 | -84.3543
0.8336 | -84.8218
0.83729 | -85.0706
0.84098 | -85.1715
0.84467 | -85.4304
0.84836 | -85.5323
0.85205 | -85.4604
0.85573 | -85.3343
0.85942 | -85.4319
0.86311 | -85.5095
0.8668 | -85.5909
0.87049 | -85.7272
0.87418 | -85.7532
0.87786 | -85.5618
0.88155 | -85.2536
0.88524 | -85.2672
0.88893 | -85.4218
0.89262 | -85.5243
0.89631 | -85.9221
0.89999 | -86.2379
0.90368 | -85.7949
0.90737 | -85.4496
0.91106 | -85.3818
0.91475 | -85.1403
0.91844 | -85.0149
0.92213 | -85.2494
0.92582 | -85.5438
0.9295 | -85.4719
0.93319 | -85.2553
0.93688 | -85.1064
0.94057 | -85.4068
0.94426 | -85.9425
0.94795 | -86.2225
0.95163 | -85.9898
0.95532 | -85.5415
0.95901 | -85.1982
0.9627 | -85.2744
0.96639 | -85.844
0.97008 | -86.579
0.97377 | -86.9978
0.97745 | -86.7588
0.98114 | -86.6418
0.98483 | -86.7746
0.98852 | -86.714
0.99221 | -86.7453
0.9959 | -86.9667
0.99958 | -86.797
1.0033 | -86.2566
1.007 | -86.0083
1.0107 | -86.0547
1.0143 | -86.3191
1.018 | -86.8274
1.0217 | -87.2939
1.0254 | -86.9632
1.0291 | -86.2052
1.0328 | -86.0007
1.0365 | -86.7941
1.0402 | -87.7574
1.0438 | -87.5521
1.0475 | -87.1033
1.0512 | -86.9066
1.0549 | -86.7175
1.0586 | -86.6893
1.0623 | -86.4791
1.066 | -86.4145
1.0697 | -86.6185
1.0734 | -87.036
1.077 | -87.3786
1.0807 | -87.4154
1.0844 | -87.2871
1.0881 | -87.3619
1.0918 | -87.77
1.0955 | -88.243
1.0992 | -88.3466
1.1029 | -88.005
1.1066 | -87.5822
1.1102 | -87.3225
1.1139 | -87.1931
1.1176 | -87.224
1.1213 | -87.4488
1.125 | -87.6946
1.1287 | -87.834
1.1324 | -87.8894
1.1361 | -87.8383
1.1397 | -87.8012
1.1434 | -87.9231
1.1471 | -88.0232
1.1508 | -87.9152
1.1545 | -87.7927
1.1582 | -87.7872
1.1619 | -87.649
1.1656 | -87.2803
1.1693 | -87.0294
1.1729 | -87.0865
1.1766 | -87.3677
1.1803 | -87.9154
1.184 | -88.6949
1.1877 | -89.2401
1.1914 | -89.303
1.1951 | -89.2491
1.1988 | -89.2376
1.2025 | -89.0832
1.2061 | -88.7844
1.2098 | -88.5816
1.2135 | -88.5737
1.2172 | -88.6347
1.2209 | -88.5724
1.2246 | -88.4248
1.2283 | -88.272
1.232 | -88.0321
1.2356 | -87.8731
1.2393 | -88.0186
1.243 | -88.3681
1.2467 | -88.7291
1.2504 | -88.9326
1.2541 | -88.9133
1.2578 | -88.8109
1.2615 | -88.7009
1.2652 | -88.5192
1.2688 | -88.3758
1.2725 | -88.5499
1.2762 | -89.0208
1.2799 | -89.4025
1.2836 | -89.5802
1.2873 | -89.6779
1.291 | -89.5868
1.2947 | -89.2869
1.2984 | -89.1069
1.302 | -89.2665
1.3057 | -89.6081
1.3094 | -89.8837
1.3131 | -90.0149
1.3168 | -90.0585
1.3205 | -90.1525
1.3242 | -90.2228
1.3279 | -90.0624
1.3316 | -89.6856
1.3352 | -89.2188
1.3389 | -88.816
1.3426 | -88.5498
1.3463 | -88.4179
1.35 | -88.4429
1.3537 | -88.5614
1.3574 | -88.7046
1.3611 | -88.8607
1.3647 | -88.9383
1.3684 | -89.0005
1.3721 | -89.1678
1.3758 | -89.3396
1.3795 | -89.4259
1.3832 | -89.42
1.3869 | -89.3722
1.3906 | -89.4194
1.3943 | -89.5252
1.3979 | -89.5193
1.4016 | -89.4163
1.4053 | -89.439
1.409 | -89.6271
1.4127 | -89.7639
1.4164 | -89.7537
1.4201 | -89.6843
1.4238 | -89.7178
1.4275 | -90.0334
1.4311 | -90.4837
1.4348 | -90.6459
1.4385 | -90.4966
1.4422 | -90.2651
1.4459 | -89.9917
1.4496 | -89.7899
1.4533 | -89.7903
1.457 | -89.9276
1.4606 | -90.1838
1.4643 | -90.5668
1.468 | -90.7549
1.4717 | -90.4566
1.4754 | -90.026
1.4791 | -89.9243
1.4828 | -90.1314
1.4865 | -90.229
1.4902 | -89.9101
1.4938 | -89.4522
1.4975 | -89.3612
1.5012 | -89.7667
1.5049 | -90.2261
1.5086 | -90.3109
1.5123 | -90.3337
1.516 | -90.7007
1.5197 | -91.2528
1.5234 | -91.5798
1.527 | -91.5579
1.5307 | -91.3851
1.5344 | -91.3276
1.5381 | -91.4184
1.5418 | -91.3056
1.5455 | -90.7728
1.5492 | -90.2526
1.5529 | -90.1795
1.5565 | -90.4365
1.5602 | -90.5679
1.5639 | -90.4815
1.5676 | -90.4212
1.5713 | -90.4687
1.575 | -90.5813
1.5787 | -90.7176
1.5824 | -90.8408
1.5861 | -90.9452
1.5897 | -90.9954
1.5934 | -90.9075
1.5971 | -90.8335
1.6008 | -90.7572
1.6045 | -90.665
1.6082 | -90.5353
1.6119 | -90.3972
1.6156 | -90.354
1.6193 | -90.4818
1.6229 | -90.7315
1.6266 | -90.9187
1.6303 | -90.8498
1.634 | -90.5953
1.6377 | -90.4017
1.6414 | -90.4345
1.6451 | -90.6682
1.6488 | -90.9109
1.6525 | -91.0039
1.6561 | -90.9672
1.6598 | -90.9513
1.6635 | -91.037
1.6672 | -91.1659
1.6709 | -91.3007
1.6746 | -91.4887
1.6783 | -91.7481
1.682 | -91.974
1.6856 | -92.0796
1.6893 | -92.1466
1.693 | -92.2624
1.6967 | -92.3897
1.7004 | -92.4256
1.7041 | -92.2771
1.7078 | -91.9419
1.7115 | -91.5584
1.7152 | -91.3489
1.7188 | -91.4387
1.7225 | -91.7339
1.7262 | -91.9659
1.7299 | -91.9423
1.7336 | -91.7383
1.7373 | -91.5637
1.741 | -91.5424
1.7447 | -91.6624
1.7484 | -91.8026
1.752 | -91.822
1.7557 | -91.6534
1.7594 | -91.3938
1.7631 | -91.234
1.7668 | -91.2995
1.7705 | -91.5545
1.7742 | -91.8598
1.7779 | -92.0916
1.7815 | -92.1753
1.7852 | -92.1065
1.7889 | -91.9741
1.7926 | -91.8679
1.7963 | -91.8223
1.8 | -91.8015
1.8037 | -91.7661
1.8074 | -91.6935
1.8111 | -91.6368
1.8147 | -91.6894
1.8184 | -91.8809
1.8221 | -92.1293
1.8258 | -92.3312
1.8295 | -92.4336
1.8332 | -92.4729
1.8369 | -92.4795
1.8406 | -92.4612
1.8443 | -92.4506
1.8479 | -92.5218
1.8516 | -92.7055
1.8553 | -92.9748
1.859 | -93.2324
1.8627 | -93.3604
1.8664 | -93.2064
1.8701 | -92.7751
1.8738 | -92.2566
1.8774 | -91.8555
1.8811 | -91.6839
1.8848 | -91.7529
1.8885 | -92.0024
1.8922 | -92.3007
1.8959 | -92.4761
1.8996 | -92.4799
1.9033 | -92.3496
1.907 | -92.1657
1.9106 | -91.9821
1.9143 | -91.8128
1.918 | -91.6486
1.9217 | -91.5405
1.9254 | -91.6035
1.9291 | -91.8769
1.9328 | -92.2458
1.9365 | -92.5715
1.9402 | -92.7673
1.9438 | -92.8426
1.9475 | -92.8697
1.9512 | -92.9314
1.9549 | -93.0396
1.9586 | -93.0891
1.9623 | -93.0156
1.966 | -92.8629
1.9697 | -92.712
1.9734 | -92.598
1.977 | -92.4981
1.9807 | -92.4101
1.9844 | -92.3705
1.9881 | -92.3796
1.9918 | -92.4189
1.9955 | -92.4459
1.9992 | -92.4672
2.0029 | -92.5359
2.0065 | -92.7161
2.0102 | -93.0172
2.0139 | -93.3545
2.0176 | -93.5666
2.0213 | -93.5503
2.025 | -93.3513
2.0287 | -93.1163
2.0324 | -92.9954
2.0361 | -93.0587
2.0397 | -93.2539
2.0434 | -93.4098
2.0471 | -93.4009
2.0508 | -93.2738
2.0545 | -93.1857
2.0582 | -93.2441
2.0619 | -93.429
2.0656 | -93.5882
2.0693 | -93.5594
2.0729 | -93.2978
2.0766 | -92.915
2.0803 | -92.5838
2.084 | -92.404
2.0877 | -92.3703
2.0914 | -92.4728
2.0951 | -92.71
2.0988 | -92.9796
2.1024 | -93.095
2.1061 | -92.9943
2.1098 | -92.8215
2.1135 | -92.6988
2.1172 | -92.6773
2.1209 | -92.7698
2.1246 | -92.9486
2.1283 | -93.0882
2.132 | -93.2121
2.1356 | -93.3074
2.1393 | -93.3981
2.143 | -93.4936
2.1467 | -93.5712
2.1504 | -93.6099
2.1541 | -93.6149
2.1578 | -93.6104
2.1615 | -93.6255
2.1652 | -93.6743
2.1688 | -93.7495
2.1725 | -93.835
2.1762 | -93.9183
2.1799 | -93.9769
2.1836 | -93.9804
2.1873 | -93.8926
2.191 | -93.6946
2.1947 | -93.4121
2.1983 | -93.1599
2.202 | -93.0648
2.2057 | -93.1793
2.2094 | -93.4627
2.2131 | -93.7831
2.2168 | -93.9808
2.2205 | -94.0175
2.2242 | -93.9589
2.2279 | -93.8887
2.2315 | -93.8555
2.2352 | -93.8472
2.2389 | -93.8274
2.2426 | -93.7644
2.2463 | -93.6694
2.25 | -93.5779
2.2537 | -93.5341
2.2574 | -93.5533
2.2611 | -93.6145
2.2647 | -93.7129
2.2684 | -93.8319
2.2721 | -93.9509
2.2758 | -94.0443
2.2795 | -94.0736
2.2832 | -94.0709
2.2869 | -94.053
2.2906 | -94.0095
2.2943 | -93.9331
2.2979 | -93.8573
2.3016 | -93.8318
2.3053 | -93.8875
2.309 | -94.0283
2.3127 | -94.1707
2.3164 | -94.2544
2.3201 | -94.2854
2.3238 | -94.2869
2.3274 | -94.3024
2.3311 | -94.291
2.3348 | -94.1665
2.3385 | -93.9126
2.3422 | -93.6165
2.3459 | -93.3998
2.3496 | -93.3303
2.3533 | -93.4364
2.357 | -93.674
2.3606 | -93.9527
2.3643 | -94.165
2.368 | -94.248
2.3717 | -94.2359
2.3754 | -94.2318
2.3791 | -94.2779
2.3828 | -94.3698
2.3865 | -94.4788
2.3902 | -94.5762
2.3938 | -94.6252
2.3975 | -94.6127
2.4012 | -94.568
2.4049 | -94.4954
2.4086 | -94.3995
2.4123 | -94.2686
2.416 | -94.0705
2.4197 | -93.8623
2.4233 | -93.7074
2.427 | -93.6537
2.4307 | -93.7386
2.4344 | -93.9609
2.4381 | -94.257
2.4418 | -94.54
2.4455 | -94.725
2.4492 | -94.762
2.4529 | -94.697
2.4565 | -94.5973
2.4602 | -94.5066
2.4639 | -94.4546
2.4676 | -94.4637
2.4713 | -94.5269
2.475 | -94.629
2.4787 | -94.7509
2.4824 | -94.8216
2.4861 | -94.8034
2.4897 | -94.7054
2.4934 | -94.5999
2.4971 | -94.5711
2.5008 | -94.6222
2.5045 | -94.7322
2.5082 | -94.8322
2.5119 | -94.8619
2.5156 | -94.7946
2.5192 | -94.6326
2.5229 | -94.4393
2.5266 | -94.2957
2.5303 | -94.2534
2.534 | -94.3155
2.5377 | -94.4426
2.5414 | -94.5733
2.5451 | -94.6673
2.5488 | -94.7444
2.5524 | -94.8166
2.5561 | -94.9025
2.5598 | -95.0167
2.5635 | -95.1438
2.5672 | -95.2574
2.5709 | -95.3434
2.5746 | -95.4046
2.5783 | -95.4663
2.582 | -95.5415
2.5856 | -95.6103
2.5893 | -95.6529
2.593 | -95.6725
2.5967 | -95.6584
2.6004 | -95.6127
2.6041 | -95.5413
2.6078 | -95.4297
2.6115 | -95.2979
2.6152 | -95.1875
2.6188 | -95.1217
2.6225 | -95.1014
2.6262 | -95.1136
2.6299 | -95.1345
2.6336 | -95.1701
2.6373 | -95.2195
2.641 | -95.2643
2.6447 | -95.3061
2.6483 | -95.3475
2.652 | -95.3768
2.6557 | -95.3823
2.6594 | -95.3669
2.6631 | -95.3653
2.6668 | -95.3775
2.6705 | -95.3938
2.6742 | -95.4122
2.6779 | -95.4005
2.6815 | -95.3581
2.6852 | -95.3012
2.6889 | -95.2526
2.6926 | -95.2413
2.6963 | -95.2613
2.7 | -95.2913
2.7037 | -95.327
2.7074 | -95.371
2.7111 | -95.4113
2.7147 | -95.4502
2.7184 | -95.4741
2.7221 | -95.4692
2.7258 | -95.4388
2.7295 | -95.3933
2.7332 | -95.3511
2.7369 | -95.3238
2.7406 | -95.3184
2.7442 | -95.3449
2.7479 | -95.3925
2.7516 | -95.4506
2.7553 | -95.5129
2.759 | -95.5713
2.7627 | -95.6136
2.7664 | -95.6366
2.7701 | -95.632
2.7738 | -95.6019
2.7774 | -95.5611
2.7811 | -95.5335
2.7848 | -95.5602
2.7885 | -95.6724
2.7922 | -95.8653
2.7959 | -96.1194
2.7996 | -96.3734
2.8033 | -96.5786
2.807 | -96.7004
2.8106 | -96.7132
2.8143 | -96.6303
2.818 | -96.502
2.8217 | -96.3795
2.8254 | -96.2911
2.8291 | -96.2562
2.8328 | -96.2427
2.8365 | -96.2227
2.8401 | -96.2004
2.8438 | -96.1842
2.8475 | -96.2027
2.8512 | -96.254
2.8549 | -96.3322
2.8586 | -96.4037
2.8623 | -96.4211
2.866 | -96.3759
2.8697 | -96.2993
2.8733 | -96.2263
2.877 | -96.1801
2.8807 | -96.1883
2.8844 | -96.2535
2.8881 | -96.3562
2.8918 | -96.4653
2.8955 | -96.5637
2.8992 | -96.6384
2.9029 | -96.6407
2.9065 | -96.5691
2.9102 | -96.4433
2.9139 | -96.3009
2.9176 | -96.1954
2.9213 | -96.1646
2.925 | -96.1846
2.9287 | -96.2228
2.9324 | -96.2453
2.9361 | -96.2455
2.9397 | -96.2464
2.9434 | -96.2615
2.9471 | -96.3129
2.9508 | -96.41
2.9545 | -96.5213
2.9582 | -96.6266
2.9619 | -96.7198
2.9656 | -96.802
2.9692 | -96.886
2.9729 | -96.971
2.9766 | -97.027
2.9803 | -97.0517
2.984 | -97.0374
2.9877 | -96.9834
2.9914 | -96.9154
2.9951 | -96.8377
2.9988 | -96.7647
3.0024 | -96.708
3.0061 | -96.6562
3.0098 | -96.6271
3.0135 | -96.6154
3.0172 | -96.6134
3.0209 | -96.6203
3.0246 | -96.6403
3.0283 | -96.677
3.032 | -96.7383
3.0356 | -96.8077
3.0393 | -96.8322
3.043 | -96.8108
3.0467 | -96.757
3.0504 | -96.6972
3.0541 | -96.6692
3.0578 | -96.681
3.0615 | -96.7445
3.0651 | -96.8298
3.0688 | -96.8989
3.0725 | -96.9361
3.0762 | -96.9402
3.0799 | -96.9377
3.0836 | -96.9426
3.0873 | -96.959
3.091 | -96.9576
3.0947 | -96.9455
3.0983 | -96.9482
3.102 | -96.9668
3.1057 | -96.9959
3.1094 | -97.0097
3.1131 | -96.9863
3.1168 | -96.922
3.1205 | -96.8238
3.1242 | -96.7266
3.1279 | -96.6441
3.1315 | -96.5756
3.1352 | -96.515
3.1389 | -96.4774
3.1426 | -96.4724
3.1463 | -96.4981
3.15 | -96.5587
3.1537 | -96.6331
3.1574 | -96.7278
3.161 | -96.8424
3.1647 | -96.972
3.1684 | -97.1069
3.1721 | -97.2039
3.1758 | -97.2341
3.1795 | -97.1923
3.1832 | -97.1001
3.1869 | -96.9803
3.1906 | -96.8631
3.1942 | -96.7611
3.1979 | -96.6689
3.2016 | -96.5913
3.2053 | -96.529
3.209 | -96.4944
3.2127 | -96.4988
3.2164 | -96.5402
3.2201 | -96.605
3.2238 | -96.7028
3.2274 | -96.8209
3.2311 | -96.9528
3.2348 | -97.0871
3.2385 | -97.2077
3.2422 | -97.2986
3.2459 | -97.3325
3.2496 | -97.3012
3.2533 | -97.2108
3.257 | -97.0775
3.2606 | -96.942
3.2643 | -96.8279
3.268 | -96.7694
3.2717 | -96.7742
3.2754 | -96.8409
3.2791 | -96.9571
3.2828 | -97.1107
3.2865 | -97.2787
3.2901 | -97.4254
3.2938 | -97.5337
3.2975 | -97.5857
3.3012 | -97.5999
3.3049 | -97.6022
3.3086 | -97.608
3.3123 | -97.6186
3.316 | -97.6292
3.3197 | -97.6492
3.3233 | -97.6846
3.327 | -97.7452
3.3307 | -97.8256
3.3344 | -97.8987
3.3381 | -97.9395
3.3418 | -97.9435
3.3455 | -97.9167
3.3492 | -97.8662
3.3529 | -97.8159
3.3565 | -97.7645
3.3602 | -97.7036
3.3639 | -97.6459
3.3676 | -97.5994
3.3713 | -97.5652
3.375 | -97.5409
3.3787 | -97.5309
3.3824 | -97.5384
3.386 | -97.5535
3.3897 | -97.5599
3.3934 | -97.5506
3.3971 | -97.5269
3.4008 | -97.4961
3.4045 | -97.4679
3.4082 | -97.4581
3.4119 | -97.4805
3.4156 | -97.5528
3.4192 | -97.6565
3.4229 | -97.7916
3.4266 | -97.9493
3.4303 | -98.1003
3.434 | -98.2326
3.4377 | -98.3273
3.4414 | -98.3843
3.4451 | -98.4171
3.4488 | -98.4148
3.4524 | -98.3693
3.4561 | -98.2906
3.4598 | -98.1828
3.4635 | -98.0575
3.4672 | -97.9379
3.4709 | -97.8268
3.4746 | -97.7363
3.4783 | -97.6678
3.4819 | -97.6306
3.4856 | -97.6262
3.4893 | -97.6531
3.493 | -97.715
3.4967 | -97.8065
3.5004 | -97.913
3.5041 | -98.0108
3.5078 | -98.0717
3.5115 | -98.0672
3.5151 | -97.999
3.5188 | -97.8722
3.5225 | -97.7251
3.5262 | -97.6017
3.5299 | -97.5273
3.5336 | -97.5105
3.5373 | -97.5503
3.541 | -97.6412
3.5447 | -97.7573
3.5483 | -97.8683
3.552 | -97.9538
3.5557 | -97.9958
3.5594 | -98.0009
3.5631 | -97.9744
3.5668 | -97.9494
3.5705 | -97.9359
3.5742 | -97.9478
3.5779 | -97.9728
3.5815 | -97.9897
3.5852 | -97.9916
3.5889 | -97.974
3.5926 | -97.9228
3.5963 | -97.8382
3.6 | -97.7389
3.6037 | -97.6556
3.6074 | -97.6078
3.611 | -97.6268
3.6147 | -97.7076
3.6184 | -97.8353
3.6221 | -97.9671
3.6258 | -98.0776
3.6295 | -98.1498
3.6332 | -98.1921
3.6369 | -98.23
3.6406 | -98.2681
3.6442 | -98.3172
3.6479 | -98.3868
3.6516 | -98.4563
3.6553 | -98.4943
3.659 | -98.4908
3.6627 | -98.4347
3.6664 | -98.3285
3.6701 | -98.1996
3.6738 | -98.0592
3.6774 | -97.9385
3.6811 | -97.8661
3.6848 | -97.8501
3.6885 | -97.8987
3.6922 | -97.9815
3.6959 | -98.076
3.6996 | -98.1469
3.7033 | -98.1828
3.7069 | -98.1921
3.7106 | -98.1919
3.7143 | -98.1969
3.718 | -98.2285
3.7217 | -98.3133
3.7254 | -98.3765
3.7291 | -98.4405
3.7328 | -98.5081
3.7365 | -98.5707
3.7401 | -98.6245
3.7438 | -98.6814
3.7475 | -98.7368
3.7512 | -98.8066
3.7549 | -98.8884
3.7586 | -98.9877
3.7623 | -99.0966
3.766 | -99.2101
3.7697 | -99.3231
3.7733 | -99.4033
3.777 | -99.4207
3.7807 | -99.3672
3.7844 | -99.2632
3.7881 | -99.1147
3.7918 | -98.9484
3.7955 | -98.7836
3.7992 | -98.6377
3.8028 | -98.5286
3.8065 | -98.4663
3.8102 | -98.452
3.8139 | -98.4869
3.8176 | -98.5576
3.8213 | -98.6472
3.825 | -98.7435
3.8287 | -98.8306
3.8324 | -98.8947
3.836 | -98.934
3.8397 | -98.9389
3.8434 | -98.9122
3.8471 | -98.8549
3.8508 | -98.7826
3.8545 | -98.7037
3.8582 | -98.6308
3.8619 | -98.5757
3.8656 | -98.5469
3.8692 | -98.5553
3.8729 | -98.6109
3.8766 | -98.7122
3.8803 | -98.8498
3.884 | -99.0117
3.8877 | -99.1741
3.8914 | -99.3039
3.8951 | -99.3926
3.8988 | -99.4343
3.9024 | -99.4296
3.9061 | -99.3728
3.9098 | -99.2817
3.9135 | -99.1719
3.9172 | -99.0587
3.9209 | -98.9516
3.9246 | -98.86
3.9283 | -98.7812
3.9319 | -98.7107
3.9356 | -98.6495
3.9393 | -98.6033
3.943 | -98.5839
3.9467 | -98.6014
3.9504 | -98.6615
3.9541 | -98.7706
3.9578 | -98.9069
3.9615 | -99.0576
3.9651 | -99.2143
3.9688 | -99.3561
3.9725 | -99.4563
3.9762 | -99.5124
3.9799 | -99.5137
3.9836 | -99.4574
3.9873 | -99.338
3.991 | -99.1853
3.9947 | -99.0236
3.9983 | -98.8745
4.002 | -98.7583
4.0057 | -98.6802
4.0094 | -98.6509
4.0131 | -98.6734
4.0168 | -98.7333
4.0205 | -98.8194
4.0242 | -98.9061
4.0278 | -98.9724
4.0315 | -99.0108
4.0352 | -99.0189
4.0389 | -98.993
4.0426 | -98.9411
4.0463 | -98.8716
4.05 | -98.8036
4.0537 | -98.7372
4.0574 | -98.6862
4.061 | -98.6726
4.0647 | -98.6997
4.0684 | -98.7717
4.0721 | -98.8695
4.0758 | -98.9718
4.0795 | -99.0656
4.0832 | -99.1279
4.0869 | -99.1545
4.0906 | -99.1455
4.0942 | -99.1126
4.0979 | -99.0809
4.1016 | -99.0725
4.1053 | -99.1048
4.109 | -99.1773
4.1127 | -99.2709
4.1164 | -99.3765
4.1201 | -99.4755
4.1237 | -99.5364
4.1274 | -99.5493
4.1311 | -99.5298
4.1348 | -99.4967
4.1385 | -99.4702
4.1422 | -99.4622
4.1459 | -99.4795
4.1496 | -99.518
4.1533 | -99.5741
4.1569 | -99.6386
4.1606 | -99.6961
4.1643 | -99.7371
4.168 | -99.7609
4.1717 | -99.7545
4.1754 | -99.7199
4.1791 | -99.6388
4.1828 | -99.5185
4.1865 | -99.3846
4.1901 | -99.2595
4.1938 | -99.142
4.1975 | -99.0555
4.2012 | -99.0098
4.2049 | -99.0074
4.2086 | -99.0458
4.2123 | -99.1091
4.216 | -99.1772
4.2197 | -99.2503
4.2233 | -99.3196
4.227 | -99.3758
4.2307 | -99.4097
4.2344 | -99.4249
4.2381 | -99.4265
4.2418 | -99.4251
4.2455 | -99.4312
4.2492 | -99.4548
4.2528 | -99.5048
4.2565 | -99.6109
4.2602 | -99.6907
4.2639 | -99.7664
4.2676 | -99.8262
4.2713 | -99.8639
4.275 | -99.8732
4.2787 | -99.8527
4.2824 | -99.8086
4.286 | -99.7474
4.2897 | -99.6936
4.2934 | -99.6523
4.2971 | -99.6282
4.3008 | -99.6255
4.3045 | -99.6252
4.3082 | -99.6256
4.3119 | -99.6192
4.3156 | -99.5919
4.3192 | -99.541
4.3229 | -99.4706
4.3266 | -99.3957
4.3303 | -99.3165
4.334 | -99.2556
4.3377 | -99.216
4.3414 | -99.2017
4.3451 | -99.2117
4.3487 | -99.2415
4.3524 | -99.2807
4.3561 | -99.3149
4.3598 | -99.3356
4.3635 | -99.3486
4.3672 | -99.3527
4.3709 | -99.352
4.3746 | -99.3593
4.3783 | -99.3593
4.3819 | -99.3542
4.3856 | -99.3595
4.3893 | -99.3735
4.393 | -99.3997
4.3967 | -99.4331
4.4004 | -99.4752
4.4041 | -99.5095
4.4078 | -99.5353
4.4115 | -99.5421
4.4151 | -99.5293
4.4188 | -99.504
4.4225 | -99.4765
4.4262 | -99.4532
4.4299 | -99.4457
4.4336 | -99.4486
4.4373 | -99.4547
4.441 | -99.4658
4.4446 | -99.4857
4.4483 | -99.5165
4.452 | -99.5524
4.4557 | -99.5848
4.4594 | -99.616
4.4631 | -99.6514
4.4668 | -99.6899
4.4705 | -99.7236
4.4742 | -99.749
4.4778 | -99.7737
4.4815 | -99.7976
4.4852 | -99.8253
4.4889 | -99.8503
4.4926 | -99.8739
4.4963 | -99.8986
4.5 | -99.9268
4.5037 | -99.9638
4.5074 | -100.0017
4.511 | -100.0254
4.5147 | -100.0357
4.5184 | -100.0403
4.5221 | -100.0269
4.5258 | -99.9942
4.5295 | -99.9434
4.5332 | -99.8844
4.5369 | -99.8408
4.5406 | -99.8228
4.5442 | -99.8412
4.5479 | -99.8927
4.5516 | -99.9781
4.5553 | -100.0874
4.559 | -100.2196
4.5627 | -100.3543
4.5664 | -100.4769
4.5701 | -100.5604
4.5737 | -100.6076
4.5774 | -100.6209
4.5811 | -100.6003
4.5848 | -100.5486
4.5885 | -100.4559
4.5922 | -100.3512
4.5959 | -100.2476
4.5996 | -100.1557
4.6033 | -100.0789
4.6069 | -100.0223
4.6106 | -99.9987
4.6143 | -99.9971
4.618 | -100.0211
4.6217 | -100.0538
4.6254 | -100.0859
4.6291 | -100.1149
4.6328 | -100.1513
4.6365 | -100.1954
4.6401 | -100.2501
4.6438 | -100.3044
4.6475 | -100.3556
4.6512 | -100.3937
4.6549 | -100.4046
4.6586 | -100.3996
4.6623 | -100.3669
4.666 | -100.3146
4.6696 | -100.2379
4.6733 | -100.1445
4.677 | -100.0471
4.6807 | -99.9503
4.6844 | -99.869
4.6881 | -99.8036
4.6918 | -99.7717
4.6955 | -99.7724
4.6992 | -99.8017
4.7028 | -99.852
4.7065 | -99.9112
4.7102 | -99.973
4.7139 | -100.0267
4.7176 | -100.0654
4.7213 | -100.0899
4.725 | -100.1067
4.7287 | -100.1224
4.7324 | -100.1344
4.736 | -100.1467
4.7397 | -100.1807
4.7434 | -100.2258
4.7471 | -100.2685
4.7508 | -100.306
4.7545 | -100.3393
4.7582 | -100.3694
4.7619 | -100.3753
4.7655 | -100.3569
4.7692 | -100.3085
4.7729 | -100.2391
4.7766 | -100.1461
4.7803 | -100.033
4.784 | -99.9254
4.7877 | -99.8859
4.7914 | -99.84
4.7951 | -99.8202
4.7987 | -99.8323
4.8024 | -99.8643
4.8061 | -99.9192
4.8098 | -99.9858
4.8135 | -100.052
4.8172 | -100.1203
4.8209 | -100.1886
4.8246 | -100.2552
4.8283 | -100.3132
4.8319 | -100.3653
4.8356 | -100.429
4.8393 | -100.4913
4.843 | -100.5492
4.8467 | -100.6021
4.8504 | -100.6584
4.8541 | -100.7151
4.8578 | -100.7605
4.8615 | -100.805
4.8651 | -100.8468
4.8688 | -100.8932
4.8725 | -100.9409
4.8762 | -100.9874
4.8799 | -101.0398
4.8836 | -101.0968
4.8873 | -101.1544
4.891 | -101.2062
4.8946 | -101.242
4.8983 | -101.2653
4.902 | -101.265
4.9057 | -101.2367
4.9094 | -101.1827
4.9131 | -101.1021
4.9168 | -101.0118
4.9205 | -100.9273
4.9242 | -100.8442
4.9278 | -100.7651
4.9315 | -100.6966
4.9352 | -100.6519
4.9389 | -100.6187
4.9426 | -100.6021
4.9463 | -100.6013
4.95 | -100.6225
4.9537 | -100.6579
4.9574 | -100.7094
4.961 | -100.7621
4.9647 | -100.813
4.9684 | -100.8577
4.9721 | -100.8861
4.9758 | -100.9007
4.9795 | -100.8989
4.9832 | -100.8914
4.9869 | -100.8723
4.9905 | -100.8412
4.9942 | -100.8
4.9979 | -100.7425
5.0016 | -100.6861
5.0053 | -100.6321
5.009 | -100.5789
5.0127 | -100.5333
5.0164 | -100.4975
5.0201 | -100.4791
5.0237 | -100.4792
5.0274 | -100.4883
5.0311 | -100.5073
5.0348 | -100.5365
5.0385 | -100.5758
5.0422 | -100.6183
5.0459 | -100.6633
5.0496 | -100.7036
5.0533 | -100.7312
5.0569 | -100.7519
5.0606 | -100.7662
5.0643 | -100.7667
5.068 | -100.7547
5.0717 | -100.7355
5.0754 | -100.7149
5.0791 | -100.6865
5.0828 | -100.6607
5.0864 | -100.6436
5.0901 | -100.642
5.0938 | -100.6687
5.0975 | -100.7152
5.1012 | -100.7729
5.1049 | -100.8461
5.1086 | -100.93
5.1123 | -101.0193
5.116 | -101.1084
5.1196 | -101.1909
5.1233 | -101.2619
5.127 | -101.3264
5.1307 | -101.3701
5.1344 | -101.3855
5.1381 | -101.3714
5.1418 | -101.342
5.1455 | -101.3019
5.1492 | -101.2462
5.1528 | -101.1711
5.1565 | -101.0993
5.1602 | -101.0363
5.1639 | -100.9777
5.1676 | -100.9166
5.1713 | -100.8664
5.175 | -100.83
5.1787 | -100.821
5.1824 | -100.8326
5.186 | -100.8599
5.1897 | -100.8998
5.1934 | -100.9534
5.1971 | -101.011
5.2008 | -101.0666
5.2045 | -101.1197
5.2082 | -101.1638
5.2119 | -101.2083
5.2155 | -101.2493
5.2192 | -101.2858
5.2229 | -101.318
5.2266 | -101.3475
5.2303 | -101.3797
5.234 | -101.4081
5.2377 | -101.4321
5.2414 | -101.4554
5.2451 | -101.4831
5.2487 | -101.5021
5.2524 | -101.5158
5.2561 | -101.5115
5.2598 | -101.497
5.2635 | -101.4746
5.2672 | -101.4503
5.2709 | -101.4192
5.2746 | -101.3998
5.2783 | -101.3952
5.2819 | -101.399
5.2856 | -101.4103
5.2893 | -101.4279
5.293 | -101.4563
5.2967 | -101.4851
5.3004 | -101.5067
5.3041 | -101.5212
5.3078 | -101.5245
5.3114 | -101.5163
5.3151 | -101.5014
5.3188 | -101.4725
5.3225 | -101.459
5.3262 | -101.4437
5.3299 | -101.4336
5.3336 | -101.4256
5.3373 | -101.4281
5.341 | -101.4371
5.3446 | -101.4464
5.3483 | -101.4604
5.352 | -101.4835
5.3557 | -101.499
5.3594 | -101.5158
5.3631 | -101.5376
5.3668 | -101.5648
5.3705 | -101.5843
5.3742 | -101.589
5.3778 | -101.5929
5.3815 | -101.5892
5.3852 | -101.5714
5.3889 | -101.5431
5.3926 | -101.5089
5.3963 | -101.4658
5.4 | -101.4187
5.4037 | -101.3708
5.4073 | -101.3273
5.411 | -101.2984
5.4147 | -101.2834
5.4184 | -101.2872
5.4221 | -101.3025
5.4258 | -101.3285
5.4295 | -101.3691
5.4332 | -101.406
5.4369 | -101.448
5.4405 | -101.4916
5.4442 | -101.5224
5.4479 | -101.5416
5.4516 | -101.5499
5.4553 | -101.5499
5.459 | -101.5457
5.4627 | -101.5303
5.4664 | -101.5073
5.4701 | -101.4805
5.4737 | -101.4521
5.4774 | -101.4225
5.4811 | -101.4069
5.4848 | -101.4064
5.4885 | -101.4152
5.4922 | -101.4344
5.4959 | -101.4675
5.4996 | -101.5051
5.5033 | -101.553
5.5069 | -101.6103
5.5106 | -101.6672
5.5143 | -101.7183
5.518 | -101.7572
5.5217 | -101.7759
5.5254 | -101.7876
5.5291 | -101.796
5.5328 | -101.7998
5.5364 | -101.7898
5.5401 | -101.773
5.5438 | -101.7555
5.5475 | -101.7426
5.5512 | -101.7314
5.5549 | -101.7252
5.5586 | -101.7246
5.5623 | -101.7321
5.566 | -101.7473
5.5696 | -101.7781
5.5733 | -101.8267
5.577 | -101.8792
5.5807 | -101.9331
5.5844 | -101.9891
5.5881 | -102.0343
5.5918 | -102.0656
5.5955 | -102.0745
5.5992 | -102.0676
5.6028 | -102.0489
5.6065 | -102.0146
5.6102 | -101.9613
5.6139 | -101.8981
5.6176 | -101.8374
5.6213 | -101.7869
5.625 | -101.747
5.6287 | -101.7154
5.6323 | -101.6886
5.636 | -101.6731
5.6397 | -101.6595
5.6434 | -101.6426
5.6471 | -101.6324
5.6508 | -101.6232
5.6545 | -101.6067
5.6582 | -101.5876
5.6619 | -101.5798
5.6655 | -101.575
5.6692 | -101.5763
5.6729 | -101.5825
5.6766 | -101.5963
5.6803 | -101.628
5.684 | -101.6688
5.6877 | -101.7233
5.6914 | -101.7879
5.6951 | -101.8456
5.6987 | -101.8947
5.7024 | -101.9397
5.7061 | -101.9637
5.7098 | -101.9592
5.7135 | -101.9314
5.7172 | -101.8818
5.7209 | -101.8192
5.7246 | -101.7511
5.7282 | -101.6828
5.7319 | -101.6195
5.7356 | -101.5686
5.7393 | -101.5332
5.743 | -101.5136
5.7467 | -101.5077
5.7504 | -101.5124
5.7541 | -101.5262
5.7578 | -101.5491
5.7614 | -101.5694
5.7651 | -101.5826
5.7688 | -101.5877
5.7725 | -101.5835
5.7762 | -101.5697
5.7799 | -101.5574
5.7836 | -101.5376
5.7873 | -101.5065
5.791 | -101.4835
5.7946 | -101.4678
5.7983 | -101.4625
5.802 | -101.4647
5.8057 | -101.465
5.8094 | -101.468
5.8131 | -101.48
5.8168 | -101.4942
5.8205 | -101.5122
5.8242 | -101.5283
5.8278 | -101.5407
5.8315 | -101.5502
5.8352 | -101.559
5.8389 | -101.569
5.8426 | -101.5819
5.8463 | -101.6021
5.85 | -101.6402
5.8537 | -101.673
5.8573 | -101.7104
5.861 | -101.7497
5.8647 | -101.7951
5.8684 | -101.8331
5.8721 | -101.8654
5.8758 | -101.8964
5.8795 | -101.9214
5.8832 | -101.9388
5.8869 | -101.9484
5.8905 | -101.9651
5.8942 | -101.9882
5.8979 | -102.0095
5.9016 | -102.0309
5.9053 | -102.0543
5.909 | -102.0793
5.9127 | -102.1107
5.9164 | -102.1435
5.9201 | -102.1735
5.9237 | -102.2032
5.9274 | -102.2203
5.9311 | -102.2294
5.9348 | -102.2318
5.9385 | -102.224
5.9422 | -102.2013
5.9459 | -102.1749
5.9496 | -102.155
5.9532 | -102.1331
5.9569 | -102.1116
5.9606 | -102.0958
5.9643 | -102.0824
5.968 | -102.0651
5.9717 | -102.0435
5.9754 | -102.0268
5.9791 | -102.0058
5.9828 | -101.9683
5.9864 | -101.9261
5.9901 | -101.8915
5.9938 | -101.8584
5.9975 | -101.8304
6.0012 | -101.8069
6.0049 | -101.7857
6.0086 | -101.7717
6.0123 | -101.7672
6.016 | -101.7696
6.0196 | -101.7882
6.0233 | -101.8112
6.027 | -101.8415
6.0307 | -101.876
6.0344 | -101.9137
6.0381 | -101.9484
6.0418 | -101.969
6.0455 | -101.9841
6.0491 | -101.9891
6.0528 | -101.9909
6.0565 | -101.9843
6.0602 | -101.9751
6.0639 | -101.9674
6.0676 | -101.9545
6.0713 | -101.9477
6.075 | -101.9469
6.0787 | -101.9485
6.0823 | -101.953
6.086 | -101.9616
6.0897 | -101.9764
6.0934 | -101.9961
6.0971 | -102.022
6.1008 | -102.0508
6.1045 | -102.0817
6.1082 | -102.1157
6.1119 | -102.1522
6.1155 | -102.2035
6.1192 | -102.262
6.1229 | -102.3167
6.1266 | -102.369
6.1303 | -102.4149
6.134 | -102.4566
6.1377 | -102.4944
6.1414 | -102.5215
6.1451 | -102.5321
6.1487 | -102.5355
6.1524 | -102.5332
6.1561 | -102.5337
6.1598 | -102.5405
6.1635 | -102.5492
6.1672 | -102.5656
6.1709 | -102.5766
6.1746 | -102.5781
6.1782 | -102.5767
6.1819 | -102.572
6.1856 | -102.5621
6.1893 | -102.5429
6.193 | -102.5207
6.1967 | -102.5048
6.2004 | -102.4997
6.2041 | -102.5023
6.2078 | -102.5142
6.2114 | -102.5335
6.2151 | -102.5696
6.2188 | -102.6135
6.2225 | -102.6662
6.2262 | -102.7199
6.2299 | -102.7674
6.2336 | -102.8205
6.2373 | -102.8731
6.241 | -102.9188
6.2446 | -102.9634
6.2483 | -102.9917
6.252 | -103.0052
6.2557 | -103.011
6.2594 | -103.0085
6.2631 | -102.9963
6.2668 | -102.9783
6.2705 | -102.9434
6.2741 | -102.9027
6.2778 | -102.8641
6.2815 | -102.8237
6.2852 | -102.782
6.2889 | -102.7383
6.2926 | -102.6929
6.2963 | -102.6482
6.3 | -102.6094
6.3037 | -102.5841
6.3073 | -102.5676
6.311 | -102.565
6.3147 | -102.581
6.3184 | -102.619
6.3221 | -102.6735
6.3258 | -102.7332
6.3295 | -102.7959
6.3332 | -102.8616
6.3369 | -102.9132
6.3405 | -102.9627
6.3442 | -102.9945
6.3479 | -103.0052
6.3516 | -103.0079
6.3553 | -103.0053
6.359 | -102.9984
6.3627 | -102.9978
6.3664 | -102.9946
6.37 | -102.9965
6.3737 | -103.0079
6.3774 | -103.0216
6.3811 | -103.0389
6.3848 | -103.0573
6.3885 | -103.069
6.3922 | -103.0775
6.3959 | -103.0773
6.3996 | -103.063
6.4032 | -103.041
6.4069 | -103.0098
6.4106 | -102.9648
6.4143 | -102.9123
6.418 | -102.8526
6.4217 | -102.7856
6.4254 | -102.7191
6.4291 | -102.6502
6.4328 | -102.5825
6.4364 | -102.5216
6.4401 | -102.4774
6.4438 | -102.4476
6.4475 | -102.4205
6.4512 | -102.4062
6.4549 | -102.4023
6.4586 | -102.4085
6.4623 | -102.426
6.466 | -102.446
6.4696 | -102.4724
6.4733 | -102.5054
6.477 | -102.5363
6.4807 | -102.5658
6.4844 | -102.5988
6.4881 | -102.6296
6.4918 | -102.6539
6.4955 | -102.6703
6.4991 | -102.6909
6.5028 | -102.7056
6.5065 | -102.7184
6.5102 | -102.7336
6.5139 | -102.7473
6.5176 | -102.7635
6.5213 | -102.7856
6.525 | -102.8105
6.5287 | -102.8381
6.5323 | -102.8692
6.536 | -102.8959
6.5397 | -102.9218
6.5434 | -102.9468
6.5471 | -102.9782
6.5508 | -103.0047
6.5545 | -103.0229
6.5582 | -103.0363
6.5619 | -103.0409
6.5655 | -103.0427
6.5692 | -103.0308
6.5729 | -103.0142
6.5766 | -102.9921
6.5803 | -102.9652
6.584 | -102.9306
6.5877 | -102.8929
6.5914 | -102.8598
6.595 | -102.8349
6.5987 | -102.8064
6.6024 | -102.7783
6.6061 | -102.7519
6.6098 | -102.7337
6.6135 | -102.7269
6.6172 | -102.7298
6.6209 | -102.7369
6.6246 | -102.7445
6.6282 | -102.7496
6.6319 | -102.7547
6.6356 | -102.7614
6.6393 | -102.7603
6.643 | -102.7582
6.6467 | -102.7457
6.6504 | -102.7277
6.6541 | -102.7106
6.6578 | -102.6923
6.6614 | -102.6698
6.6651 | -102.644
6.6688 | -102.6237
6.6725 | -102.5999
6.6762 | -102.5854
6.6799 | -102.5712
6.6836 | -102.5595
6.6873 | -102.5634
6.6909 | -102.5704
6.6946 | -102.5778
6.6983 | -102.5904
6.702 | -102.6128
6.7057 | -102.6343
6.7094 | -102.6498
6.7131 | -102.6651
6.7168 | -102.6779
6.7205 | -102.6869
6.7241 | -102.6948
6.7278 | -102.7055
6.7315 | -102.7186
6.7352 | -102.7283
6.7389 | -102.741
6.7426 | -102.7451
6.7463 | -102.7489
6.75 | -102.7535
6.7537 | -102.7452
6.7573 | -102.7261
6.761 | -102.7044
6.7647 | -102.6845
6.7684 | -102.6627
6.7721 | -102.646
6.7758 | -102.628
6.7795 | -102.6199
6.7832 | -102.617
6.7869 | -102.6131
6.7905 | -102.6185
6.7942 | -102.6319
6.7979 | -102.6513
6.8016 | -102.6714
6.8053 | -102.696
6.809 | -102.7191
6.8127 | -102.7365
6.8164 | -102.7529
6.82 | -102.7632
6.8237 | -102.7669
6.8274 | -102.7655
6.8311 | -102.7549
6.8348 | -102.7424
6.8385 | -102.732
6.8422 | -102.7265
6.8459 | -102.7265
6.8496 | -102.7301
6.8532 | -102.7445
6.8569 | -102.7683
6.8606 | -102.7969
6.8643 | -102.8256
6.868 | -102.8614
6.8717 | -102.8963
6.8754 | -102.9351
6.8791 | -102.9756
6.8828 | -103.0074
6.8864 | -103.0299
6.8901 | -103.0488
6.8938 | -103.0548
6.8975 | -103.0485
6.9012 | -103.0341
6.9049 | -103.0081
6.9086 | -102.9741
6.9123 | -102.9404
6.9159 | -102.919
6.9196 | -102.8924
6.9233 | -102.8719
6.927 | -102.8575
6.9307 | -102.846
6.9344 | -102.8431
6.9381 | -102.8492
6.9418 | -102.8572
6.9455 | -102.8697
6.9491 | -102.886
6.9528 | -102.9089
6.9565 | -102.9335
6.9602 | -102.963
6.9639 | -102.9968
6.9676 | -103.0308
6.9713 | -103.0634
6.975 | -103.0959
6.9787 | -103.1258
6.9823 | -103.1417
6.986 | -103.1574
6.9897 | -103.1688
6.9934 | -103.1694
6.9971 | -103.1713
7.0008 | -103.1714
7.0045 | -103.1681
7.0082 | -103.1661
7.0118 | -103.1666
7.0155 | -103.1619
7.0192 | -103.1567
7.0229 | -103.1553
7.0266 | -103.1621
7.0303 | -103.1735
7.034 | -103.1803
7.0377 | -103.1917
7.0414 | -103.2074
7.045 | -103.2277
7.0487 | -103.2581
7.0524 | -103.2948
7.0561 | -103.3318
7.0598 | -103.3637
7.0635 | -103.3965
7.0672 | -103.423
7.0709 | -103.448
7.0746 | -103.4666
7.0782 | -103.4801
7.0819 | -103.4876
7.0856 | -103.4887
7.0893 | -103.4886
7.093 | -103.4769
7.0967 | -103.4606
7.1004 | -103.4391
7.1041 | -103.4148
7.1078 | -103.3888
7.1114 | -103.3643
7.1151 | -103.3344
7.1188 | -103.3102
7.1225 | -103.2847
7.1262 | -103.26
7.1299 | -103.2443
7.1336 | -103.233
7.1373 | -103.2256
7.1409 | -103.2226
7.1446 | -103.2227
7.1483 | -103.2258
7.152 | -103.2301
7.1557 | -103.2335
7.1594 | -103.2384
7.1631 | -103.2462
7.1668 | -103.2598
7.1705 | -103.2723
7.1741 | -103.2809
7.1778 | -103.2802
7.1815 | -103.2804
7.1852 | -103.2779
7.1889 | -103.2711
7.1926 | -103.2626
7.1963 | -103.2583
7.2 | -103.2512
7.2037 | -103.239
7.2073 | -103.2312
7.211 | -103.2238
7.2147 | -103.2173
7.2184 | -103.2132
7.2221 | -103.2154
7.2258 | -103.2158
7.2295 | -103.2167
7.2332 | -103.2188
7.2368 | -103.2213
7.2405 | -103.225
7.2442 | -103.2308
7.2479 | -103.2443
7.2516 | -103.2562
7.2553 | -103.2594
7.259 | -103.2608
7.2627 | -103.2686
7.2664 | -103.2754
7.27 | -103.2782
7.2737 | -103.2838
7.2774 | -103.29
7.2811 | -103.2962
7.2848 | -103.3045
7.2885 | -103.309
7.2922 | -103.3187
7.2959 | -103.336
7.2996 | -103.3578
7.3032 | -103.3806
7.3069 | -103.4147
7.3106 | -103.4521
7.3143 | -103.4898
7.318 | -103.532
7.3217 | -103.5744
7.3254 | -103.6178
7.3291 | -103.6591
7.3327 | -103.6943
7.3364 | -103.7297
7.3401 | -103.7535
7.3438 | -103.7665
7.3475 | -103.7768
7.3512 | -103.7846
7.3549 | -103.7891
7.3586 | -103.7974
7.3623 | -103.8126
7.3659 | -103.8262
7.3696 | -103.8332
7.3733 | -103.8442
7.377 | -103.8563
7.3807 | -103.8605
7.3844 | -103.8675
7.3881 | -103.8709
7.3918 | -103.8632
7.3955 | -103.8492
7.3991 | -103.8221
7.4028 | -103.7896
7.4065 | -103.7476
7.4102 | -103.7033
7.4139 | -103.6557
7.4176 | -103.6055
7.4213 | -103.5541
7.425 | -103.5025
7.4287 | -103.4592
7.4323 | -103.4206
7.436 | -103.3913
7.4397 | -103.3767
7.4434 | -103.3733
7.4471 | -103.4103
7.4508 | -103.4278
7.4545 | -103.4562
7.4582 | -103.4881
7.4618 | -103.5219
7.4655 | -103.5558
7.4692 | -103.5874
7.4729 | -103.6189
7.4766 | -103.6503
7.4803 | -103.6744
7.484 | -103.694
7.4877 | -103.7103
7.4914 | -103.7206
7.495 | -103.723
7.4987 | -103.7197
7.5024 | -103.7183
7.5061 | -103.7183
7.5098 | -103.7182
7.5135 | -103.717
7.5172 | -103.7198
7.5209 | -103.7144
7.5246 | -103.7082
7.5282 | -103.7064
7.5319 | -103.7035
7.5356 | -103.6995
7.5393 | -103.6842
7.543 | -103.6608
7.5467 | -103.6325
7.5504 | -103.6011
7.5541 | -103.5703
7.5577 | -103.5419
7.5614 | -103.5177
7.5651 | -103.4935
7.5688 | -103.4757
7.5725 | -103.4687
7.5762 | -103.4719
7.5799 | -103.4888
7.5836 | -103.5161
7.5873 | -103.5534
7.5909 | -103.5987
7.5946 | -103.6531
7.5983 | -103.713
7.602 | -103.7753
7.6057 | -103.8421
7.6094 | -103.9039
7.6131 | -103.9653
7.6168 | -104.0155
7.6205 | -104.0517
7.6241 | -104.0794
7.6278 | -104.1006
7.6315 | -104.1166
7.6352 | -104.1331
7.6389 | -104.1346
7.6426 | -104.1197
7.6463 | -104.099
7.65 | -104.0733
7.6536 | -104.0575
7.6573 | -104.0407
7.661 | -104.0232
7.6647 | -104.0027
7.6684 | -103.9797
7.6721 | -103.9594
7.6758 | -103.9391
7.6795 | -103.9172
7.6832 | -103.8924
7.6868 | -103.8697
7.6905 | -103.8439
7.6942 | -103.8086
7.6979 | -103.7734
7.7016 | -103.7387
7.7053 | -103.7016
7.709 | -103.6678
7.7127 | -103.6393
7.7164 | -103.6178
7.72 | -103.5965
7.7237 | -103.5819
7.7274 | -103.5726
7.7311 | -103.5706
7.7348 | -103.5763
7.7385 | -103.5881
7.7422 | -103.6023
7.7459 | -103.6186
7.7496 | -103.6389
7.7532 | -103.6663
7.7569 | -103.6979
7.7606 | -103.7307
7.7643 | -103.7652
7.768 | -103.7946
7.7717 | -103.8259
7.7754 | -103.8593
7.7791 | -103.8951
7.7827 | -103.9337
7.7864 | -103.9661
7.7901 | -104.0034
7.7938 | -104.0441
7.7975 | -104.0727
7.8012 | -104.0913
7.8049 | -104.1063
7.8086 | -104.1202
7.8123 | -104.1288
7.8159 | -104.128
7.8196 | -104.1148
7.8233 | -104.0946
7.827 | -104.075
7.8307 | -104.0534
7.8344 | -104.0346
7.8381 | -104.0212
7.8418 | -104.0081
7.8455 | -103.9949
7.8491 | -103.98
7.8528 | -103.974
7.8565 | -103.97
7.8602 | -103.9626
7.8639 | -103.9536
7.8676 | -103.939
7.8713 | -103.9228
7.875 | -103.8984
7.8786 | -103.8714
7.8823 | -103.8397
7.886 | -103.805
7.8897 | -103.7688
7.8934 | -103.7302
7.8971 | -103.6906
7.9008 | -103.6542
7.9045 | -103.6205
7.9082 | -103.5926
7.9118 | -103.5789
7.9155 | -103.5743
7.9192 | -103.5811
7.9229 | -103.5963
7.9266 | -103.6175
7.9303 | -103.65
7.934 | -103.6949
7.9377 | -103.7448
7.9414 | -103.7886
7.945 | -103.8338
7.9487 | -103.8851
7.9524 | -103.9314
7.9561 | -103.9747
7.9598 | -104.0127
7.9635 | -104.045
7.9672 | -104.0643
7.9709 | -104.0756
7.9745 | -104.0823
7.9782 | -104.0824
7.9819 | -104.0654
7.9856 | -104.0611
7.9893 | -104.0504
7.993 | -104.0433
7.9967 | -104.0358
8.0004 | -104.0306
8.0041 | -104.0305
8.0077 | -104.035
8.0114 | -104.0395
8.0151 | -104.0521
8.0188 | -104.0728
8.0225 | -104.0979
8.0262 | -104.1241
8.0299 | -104.15
8.0336 | -104.1727
8.0373 | -104.1935
8.0409 | -104.2143
8.0446 | -104.2248
8.0483 | -104.2268
8.052 | -104.2253
8.0557 | -104.2171
8.0594 | -104.2121
8.0631 | -104.2088
8.0668 | -104.1988
8.0704 | -104.1891
8.0741 | -104.1827
8.0778 | -104.1831
8.0815 | -104.1902
8.0852 | -104.197
8.0889 | -104.2108
8.0926 | -104.2269
8.0963 | -104.2401
8.1 | -104.2535
8.1036 | -104.2692
8.1073 | -104.2832
8.111 | -104.2939
8.1147 | -104.3004
8.1184 | -104.3034
8.1221 | -104.3014
8.1258 | -104.293
8.1295 | -104.2759
8.1332 | -104.2522
8.1368 | -104.2179
8.1405 | -104.1819
8.1442 | -104.142
8.1479 | -104.1027
8.1516 | -104.0667
8.1553 | -104.0381
8.159 | -104.0137
8.1627 | -103.9976
8.1664 | -103.997
8.17 | -104.0057
8.1737 | -104.0197
8.1774 | -104.0463
8.1811 | -104.0754
8.1848 | -104.104
8.1885 | -104.1321
8.1922 | -104.1614
8.1959 | -104.1861
8.1995 | -104.1976
8.2032 | -104.1981
8.2069 | -104.1929
8.2106 | -104.1766
8.2143 | -104.1512
8.218 | -104.117
8.2217 | -104.0695
8.2254 | -104.0151
8.2291 | -103.963
8.2327 | -103.9098
8.2364 | -103.8568
8.2401 | -103.8067
8.2438 | -103.7637
8.2475 | -103.7282
8.2512 | -103.7062
8.2549 | -103.6961
8.2586 | -103.6945
8.2623 | -103.6969
8.2659 | -103.7071
8.2696 | -103.7271
8.2733 | -103.744
8.277 | -103.7612
8.2807 | -103.7782
8.2844 | -103.788
8.2881 | -103.7955
8.2918 | -103.7974
8.2954 | -103.7972
8.2991 | -103.7926
8.3028 | -103.7883
8.3065 | -103.7899
8.3102 | -103.7992
8.3139 | -103.8118
8.3176 | -103.8311
8.3213 | -103.8585
8.325 | -103.8899
8.3286 | -103.9278
8.3323 | -103.9723
8.336 | -104.0168
8.3397 | -104.0707
8.3434 | -104.1296
8.3471 | -104.1811
8.3508 | -104.2282
8.3545 | -104.2706
8.3582 | -104.3054
8.3618 | -104.337
8.3655 | -104.36
8.3692 | -104.3774
8.3729 | -104.3862
8.3766 | -104.3922
8.3803 | -104.3896
8.384 | -104.3855
8.3877 | -104.3813
8.3913 | -104.3772
8.395 | -104.3787
8.3987 | -104.3751
8.4024 | -104.3737
8.4061 | -104.3797
8.4098 | -104.3888
8.4135 | -104.3995
8.4172 | -104.4062
8.4209 | -104.4138
8.4245 | -104.4243
8.4282 | -104.4374
8.4319 | -104.4498
8.4356 | -104.4606
8.4393 | -104.4727
8.443 | -104.4896
8.4467 | -104.507
8.4504 | -104.5216
8.4541 | -104.5351
8.4577 | -104.5526
8.4614 | -104.5624
8.4651 | -104.5685
8.4688 | -104.5743
8.4725 | -104.572
8.4762 | -104.5657
8.4799 | -104.551
8.4836 | -104.5293
8.4873 | -104.5067
8.4909 | -104.476
8.4946 | -104.4401
8.4983 | -104.4015
8.502 | -104.3703
8.5057 | -104.3442
8.5094 | -104.3186
8.5131 | -104.3081
8.5168 | -104.2894
8.5204 | -104.2814
8.5241 | -104.2833
8.5278 | -104.2878
8.5315 | -104.2957
8.5352 | -104.3043
8.5389 | -104.3197
8.5426 | -104.3388
8.5463 | -104.3573
8.55 | -104.381
8.5536 | -104.4007
8.5573 | -104.4166
8.561 | -104.433
8.5647 | -104.4458
8.5684 | -104.4511
8.5721 | -104.458
8.5758 | -104.4672
8.5795 | -104.473
8.5832 | -104.4751
8.5868 | -104.4798
8.5905 | -104.4844
8.5942 | -104.4812
8.5979 | -104.4762
8.6016 | -104.4699
8.6053 | -104.4655
8.609 | -104.4609
8.6127 | -104.4514
8.6163 | -104.4432
8.62 | -104.4406
8.6237 | -104.4363
8.6274 | -104.4305
8.6311 | -104.4215
8.6348 | -104.4195
8.6385 | -104.4182
8.6422 | -104.4168
8.6459 | -104.4178
8.6495 | -104.4177
8.6532 | -104.4183
8.6569 | -104.4233
8.6606 | -104.4291
8.6643 | -104.4388
8.668 | -104.447
8.6717 | -104.4554
8.6754 | -104.4632
8.6791 | -104.468
8.6827 | -104.4729
8.6864 | -104.4804
8.6901 | -104.4875
8.6938 | -104.4939
8.6975 | -104.5062
8.7012 | -104.5196
8.7049 | -104.5275
8.7086 | -104.5358
8.7122 | -104.5445
8.7159 | -104.5447
8.7196 | -104.5421
8.7233 | -104.5412
8.727 | -104.5347
8.7307 | -104.5285
8.7344 | -104.5174
8.7381 | -104.5
8.7418 | -104.484
8.7454 | -104.4698
8.7491 | -104.4524
8.7528 | -104.4387
8.7565 | -104.4283
8.7602 | -104.4262
8.7639 | -104.4261
8.7676 | -104.4334
8.7713 | -104.4427
8.775 | -104.4497
8.7786 | -104.458
8.7823 | -104.4693
8.786 | -104.4821
8.7897 | -104.4967
8.7934 | -104.512
8.7971 | -104.5273
8.8008 | -104.5376
8.8045 | -104.5428
8.8082 | -104.5442
8.8118 | -104.5434
8.8155 | -104.5362
8.8192 | -104.5265
8.8229 | -104.51
8.8266 | -104.4864
8.8303 | -104.4613
8.834 | -104.4371
8.8377 | -104.4108
8.8413 | -104.3854
8.845 | -104.3631
8.8487 | -104.3386
8.8524 | -104.3185
8.8561 | -104.3022
8.8598 | -104.2826
8.8635 | -104.2709
8.8672 | -104.2666
8.8709 | -104.2657
8.8745 | -104.2668
8.8782 | -104.2745
8.8819 | -104.2852
8.8856 | -104.2967
8.8893 | -104.3074
8.893 | -104.3242
8.8967 | -104.3446
8.9004 | -104.369
8.9041 | -104.4001
8.9077 | -104.4317
8.9114 | -104.4628
8.9151 | -104.4964
8.9188 | -104.5334
8.9225 | -104.5691
8.9262 | -104.6082
8.9299 | -104.6492
8.9336 | -104.691
8.9372 | -104.7264
8.9409 | -104.7586
8.9446 | -104.791
8.9483 | -104.8179
8.952 | -104.8425
8.9557 | -104.8605
8.9594 | -104.8709
8.9631 | -104.8785
8.9668 | -104.8841
8.9704 | -104.8819
8.9741 | -104.8765
8.9778 | -104.8629
8.9815 | -104.8389
8.9852 | -104.8115
8.9889 | -104.7838
8.9926 | -104.7553
8.9963 | -104.7237
9 | -104.6939
9.0036 | -104.6686
9.0073 | -104.6462
9.011 | -104.6293
9.0147 | -104.6188
9.0184 | -104.6117
9.0221 | -104.6049
9.0258 | -104.6062
9.0295 | -104.613
9.0331 | -104.6261
9.0368 | -104.6493
9.0405 | -104.6771
9.0442 | -104.7136
9.0479 | -104.7399
9.0516 | -104.7638
9.0553 | -104.785
9.059 | -104.8046
9.0627 | -104.8244
9.0663 | -104.8366
9.07 | -104.8463
9.0737 | -104.8542
9.0774 | -104.8569
9.0811 | -104.8508
9.0848 | -104.8395
9.0885 | -104.8282
9.0922 | -104.8141
9.0959 | -104.7971
9.0995 | -104.7819
9.1032 | -104.7682
9.1069 | -104.7533
9.1106 | -104.7467
9.1143 | -104.7413
9.118 | -104.7408
9.1217 | -104.7381
9.1254 | -104.7375
9.1291 | -104.7428
9.1327 | -104.7493
9.1364 | -104.7589
9.1401 | -104.7697
9.1438 | -104.7787
9.1475 | -104.7866
9.1512 | -104.798
9.1549 | -104.808
9.1586 | -104.8098
9.1622 | -104.8157
9.1659 | -104.8192
9.1696 | -104.819
9.1733 | -104.8189
9.177 | -104.8192
9.1807 | -104.8196
9.1844 | -104.8232
9.1881 | -104.8252
9.1918 | -104.8261
9.1954 | -104.8255
9.1991 | -104.8224
9.2028 | -104.8213
9.2065 | -104.8189
9.2102 | -104.8109
9.2139 | -104.7986
9.2176 | -104.7863
9.2213 | -104.7737
9.225 | -104.7538
9.2286 | -104.7304
9.2323 | -104.7072
9.236 | -104.6803
9.2397 | -104.6554
9.2434 | -104.6334
9.2471 | -104.61
9.2508 | -104.5913
9.2545 | -104.5715
9.2581 | -104.557
9.2618 | -104.5459
9.2655 | -104.5418
9.2692 | -104.5463
9.2729 | -104.5491
9.2766 | -104.5566
9.2803 | -104.5746
9.284 | -104.5971
9.2877 | -104.6194
9.2913 | -104.6468
9.295 | -104.6755
9.2987 | -104.703
9.3024 | -104.7295
9.3061 | -104.7579
9.3098 | -104.7841
9.3135 | -104.8057
9.3172 | -104.8269
9.3209 | -104.8445
9.3245 | -104.8638
9.3282 | -104.8821
9.3319 | -104.8981
9.3356 | -104.9129
9.3393 | -104.9301
9.343 | -104.9513
9.3467 | -104.9766
9.3504 | -105.0047
9.3541 | -105.0358
9.3577 | -105.065
9.3614 | -105.1016
9.3651 | -105.1426
9.3688 | -105.1781
9.3725 | -105.2165
9.3762 | -105.2519
9.3799 | -105.2828
9.3836 | -105.3116
9.3872 | -105.3394
9.3909 | -105.3644
9.3946 | -105.3874
9.3983 | -105.4024
9.402 | -105.41
9.4057 | -105.4116
9.4094 | -105.4173
9.4131 | -105.4189
9.4168 | -105.4139
9.4204 | -105.4052
9.4241 | -105.3935
9.4278 | -105.3768
9.4315 | -105.3599
9.4352 | -105.3388
9.4389 | -105.3126
9.4426 | -105.284
9.4463 | -105.2534
9.45 | -105.223
9.4536 | -105.1883
9.4573 | -105.1508
9.461 | -105.1136
9.4647 | -105.0733
9.4684 | -105.0315
9.4721 | -104.9959
9.4758 | -104.9587
9.4795 | -104.9192
9.4831 | -104.8864
9.4868 | -104.8606
9.4905 | -104.8396
9.4942 | -104.8242
9.4979 | -104.8073
9.5016 | -104.7935
9.5053 | -104.7919
9.509 | -104.7926
9.5127 | -104.7935
9.5163 | -104.8044
9.52 | -104.8169
9.5237 | -104.8263
9.5274 | -104.8374
9.5311 | -104.8521
9.5348 | -104.8657
9.5385 | -104.8741
9.5422 | -104.8817
9.5459 | -104.8856
9.5495 | -104.8889
9.5532 | -104.8871
9.5569 | -104.8843
9.5606 | -104.8884
9.5643 | -104.889
9.568 | -104.8857
9.5717 | -104.8854
9.5754 | -104.8921
9.579 | -104.8994
9.5827 | -104.9057
9.5864 | -104.9129
9.5901 | -104.9221
9.5938 | -104.9293
9.5975 | -104.9417
9.6012 | -104.9556
9.6049 | -104.9709
9.6086 | -104.9814
9.6122 | -104.9883
9.6159 | -104.9965
9.6196 | -105.0061
9.6233 | -105.0095
9.627 | -105.0103
9.6307 | -105.0087
9.6344 | -105.0083
9.6381 | -105.007
9.6418 | -105.0028
9.6454 | -104.9974
9.6491 | -104.9956
9.6528 | -104.9963
9.6565 | -104.9894
9.6602 | -104.9842
9.6639 | -104.9815
9.6676 | -104.9804
9.6713 | -104.9792
9.675 | -104.981
9.6786 | -104.9866
9.6823 | -104.9974
9.686 | -105.0111
9.6897 | -105.0232
9.6934 | -105.0416
9.6971 | -105.0632
9.7008 | -105.0818
9.7045 | -105.1044
9.7081 | -105.1296
9.7118 | -105.1494
9.7155 | -105.1729
9.7192 | -105.1955
9.7229 | -105.2133
9.7266 | -105.2333
9.7303 | -105.2499
9.734 | -105.2654
9.7377 | -105.2856
9.7413 | -105.3012
9.745 | -105.3146
9.7487 | -105.3278
9.7524 | -105.3421
9.7561 | -105.3561
9.7598 | -105.3648
9.7635 | -105.3686
9.7672 | -105.3698
9.7709 | -105.3695
9.7745 | -105.3671
9.7782 | -105.3609
9.7819 | -105.3535
9.7856 | -105.346
9.7893 | -105.3316
9.793 | -105.3168
9.7967 | -105.3058
9.8004 | -105.2961
9.804 | -105.283
9.8077 | -105.2709
9.8114 | -105.2591
9.8151 | -105.2411
9.8188 | -105.2254
9.8225 | -105.207
9.8262 | -105.1873
9.8299 | -105.1693
9.8336 | -105.1546
9.8372 | -105.1351
9.8409 | -105.1141
9.8446 | -105.0941
9.8483 | -105.0747
9.852 | -105.0572
9.8557 | -105.0401
9.8594 | -105.0226
9.8631 | -105.007
9.8668 | -104.9972
9.8704 | -104.986
9.8741 | -104.9798
9.8778 | -104.9734
9.8815 | -104.969
9.8852 | -104.9657
9.8889 | -104.9638
9.8926 | -104.9669
9.8963 | -104.975
9.8999 | -104.9803
9.9036 | -104.9878
9.9073 | -104.9965
9.911 | -105.0051
9.9147 | -105.0091
9.9184 | -105.0114
9.9221 | -105.0104
9.9258 | -105.0092
9.9295 | -105.0077
9.9331 | -105.0076
9.9368 | -105.0062
9.9405 | -105.0042
9.9442 | -105.0046
9.9479 | -105.0057
9.9516 | -105.0009
9.9553 | -104.9976
9.959 | -104.9999
9.9627 | -104.9997
9.9663 | -104.9978
9.97 | -104.9936
9.9737 | -104.9889
9.9774 | -104.9809
9.9811 | -104.9766
9.9848 | -104.9754
9.9885 | -104.9741
9.9922 | -104.9739
9.9959 | -104.969
9.9995 | -104.9713
10.0032 | -104.9701
10.0069 | -104.9685
10.0106 | -104.9687
10.0143 | -104.9728
10.018 | -104.9782
10.0217 | -104.9811
10.0254 | -104.986
10.029 | -104.9947
10.0327 | -105.0049
10.0364 | -105.0166
10.0401 | -105.0269
10.0438 | -105.0431
10.0475 | -105.0591
10.0512 | -105.0776
10.0549 | -105.0954
10.0586 | -105.1119
10.0622 | -105.1259
10.0659 | -105.1396
10.0696 | -105.152
10.0733 | -105.1585
10.077 | -105.1634
10.0807 | -105.1643
10.0844 | -105.1607
10.0881 | -105.1557
10.0918 | -105.1534
10.0954 | -105.1502
10.0991 | -105.1394
10.1028 | -105.1273
10.1065 | -105.1156
10.1102 | -105.1082
10.1139 | -105.1016
10.1176 | -105.0929
10.1213 | -105.0858
10.1249 | -105.0846
10.1286 | -105.0823
10.1323 | -105.078
10.136 | -105.075
10.1397 | -105.0753
10.1434 | -105.0839
10.1471 | -105.0901
10.1508 | -105.1
10.1545 | -105.1135
10.1581 | -105.1257
10.1618 | -105.1364
10.1655 | -105.1504
10.1692 | -105.1685
10.1729 | -105.1841
10.1766 | -105.1999
10.1803 | -105.2193
10.184 | -105.2367
10.1877 | -105.2564
10.1913 | -105.2758
10.195 | -105.2958
10.1987 | -105.3139
10.2024 | -105.3356
10.2061 | -105.358
10.2098 | -105.3765
10.2135 | -105.3917
10.2172 | -105.4045
10.2208 | -105.4164
10.2245 | -105.4287
10.2282 | -105.4377
10.2319 | -105.4479
10.2356 | -105.456
10.2393 | -105.4601
10.243 | -105.4677
10.2467 | -105.4734
10.2504 | -105.4742
10.254 | -105.4696
10.2577 | -105.464
10.2614 | -105.4562
10.2651 | -105.4503
10.2688 | -105.4442
10.2725 | -105.4387
10.2762 | -105.4344
10.2799 | -105.4307
10.2836 | -105.4305
10.2872 | -105.4317
10.2909 | -105.4336
10.2946 | -105.4324
10.2983 | -105.4378
10.302 | -105.4472
10.3057 | -105.4551
10.3094 | -105.4667
10.3131 | -105.4825
10.3168 | -105.4917
10.3204 | -105.4998
10.3241 | -105.5088
10.3278 | -105.5165
10.3315 | -105.5176
10.3352 | -105.5174
10.3389 | -105.5112
10.3426 | -105.5021
10.3463 | -105.4876
10.3499 | -105.47
10.3536 | -105.4483
10.3573 | -105.4244
10.361 | -105.4021
10.3647 | -105.3781
10.3684 | -105.3507
10.3721 | -105.3247
10.3758 | -105.2976
10.3795 | -105.2746
10.3831 | -105.2527
10.3868 | -105.2337
10.3905 | -105.2219
10.3942 | -105.2142
10.3979 | -105.2158
10.4016 | -105.2241
10.4053 | -105.2405
10.409 | -105.261
10.4127 | -105.2861
10.4163 | -105.3165
10.42 | -105.3508
10.4237 | -105.3863
10.4274 | -105.4243
10.4311 | -105.4655
10.4348 | -105.5078
10.4385 | -105.5541
10.4422 | -105.5984
10.4458 | -105.6426
10.4495 | -105.6809
10.4532 | -105.7155
10.4569 | -105.7505
10.4606 | -105.7797
10.4643 | -105.8067
10.468 | -105.83
10.4717 | -105.8502
10.4754 | -105.8632
10.479 | -105.8692
10.4827 | -105.8721
10.4864 | -105.8703
10.4901 | -105.8659
10.4938 | -105.8589
10.4975 | -105.8514
10.5012 | -105.8396
10.5049 | -105.8234
10.5086 | -105.8034
10.5122 | -105.7799
10.5159 | -105.756
10.5196 | -105.7331
10.5233 | -105.7078
10.527 | -105.6808
10.5307 | -105.6539
10.5344 | -105.6274
10.5381 | -105.6032
10.5417 | -105.58
10.5454 | -105.555
10.5491 | -105.5332
10.5528 | -105.5137
10.5565 | -105.4901
10.5602 | -105.4689
10.5639 | -105.4513
10.5676 | -105.4389
10.5713 | -105.4267
10.5749 | -105.4172
10.5786 | -105.413
10.5823 | -105.4137
10.586 | -105.4182
10.5897 | -105.4241
10.5934 | -105.432
10.5971 | -105.4433
10.6008 | -105.4549
10.6045 | -105.4693
10.6081 | -105.4815
10.6118 | -105.4947
10.6155 | -105.5073
10.6192 | -105.5247
10.6229 | -105.54
10.6266 | -105.5571
10.6303 | -105.5732
10.634 | -105.5875
10.6377 | -105.6019
10.6413 | -105.6134
10.645 | -105.6283
10.6487 | -105.6396
10.6524 | -105.6485
10.6561 | -105.657
10.6598 | -105.6632
10.6635 | -105.6707
10.6672 | -105.6777
10.6708 | -105.6793
10.6745 | -105.6782
10.6782 | -105.6787
10.6819 | -105.6769
10.6856 | -105.6714
10.6893 | -105.667
10.693 | -105.6604
10.6967 | -105.6546
10.7004 | -105.6513
10.704 | -105.6457
10.7077 | -105.6418
10.7114 | -105.645
10.7151 | -105.6461
10.7188 | -105.6432
10.7225 | -105.6372
10.7262 | -105.6377
10.7299 | -105.6383
10.7336 | -105.6331
10.7372 | -105.622
10.7409 | -105.6089
10.7446 | -105.5921
10.7483 | -105.5759
10.752 | -105.5552
10.7557 | -105.534
10.7594 | -105.5084
10.7631 | -105.4829
10.7667 | -105.4568
10.7704 | -105.4299
10.7741 | -105.4018
10.7778 | -105.3762
10.7815 | -105.3512
10.7852 | -105.3325
10.7889 | -105.3155
10.7926 | -105.3018
10.7963 | -105.2898
10.7999 | -105.2814
10.8036 | -105.2819
10.8073 | -105.2883
10.811 | -105.2989
10.8147 | -105.31
10.8184 | -105.327
10.8221 | -105.3478
10.8258 | -105.3716
10.8295 | -105.402
10.8331 | -105.4339
10.8368 | -105.4678
10.8405 | -105.5052
10.8442 | -105.5468
10.8479 | -105.5868
10.8516 | -105.6257
10.8553 | -105.6626
10.859 | -105.7031
10.8626 | -105.7419
10.8663 | -105.7755
10.87 | -105.8064
10.8737 | -105.8349
10.8774 | -105.8645
10.8811 | -105.8867
10.8848 | -105.8945
10.8885 | -105.8969
10.8922 | -105.8998
10.8958 | -105.9023
10.8995 | -105.8959
10.9032 | -105.8865
10.9069 | -105.8721
10.9106 | -105.8532
10.9143 | -105.8303
10.918 | -105.8073
10.9217 | -105.7839
10.9254 | -105.7607
10.929 | -105.7333
10.9327 | -105.7059
10.9364 | -105.6767
10.9401 | -105.649
10.9438 | -105.6244
10.9475 | -105.6009
10.9512 | -105.5779
10.9549 | -105.5551
10.9586 | -105.5328
10.9622 | -105.5092
10.9659 | -105.4859
10.9696 | -105.4665
10.9733 | -105.4496
10.977 | -105.4309
10.9807 | -105.4092
10.9844 | -105.3876
10.9881 | -105.3663
10.9917 | -105.3497
10.9954 | -105.3353
10.9991 | -105.3217
11.0028 | -105.3105
11.0065 | -105.3014
11.0102 | -105.2966
11.0139 | -105.2911
11.0176 | -105.2845
11.0213 | -105.2794
11.0249 | -105.2748
11.0286 | -105.2693
11.0323 | -105.2622
11.036 | -105.2537
11.0397 | -105.2499
11.0434 | -105.2495
11.0471 | -105.2464
11.0508 | -105.2402
11.0545 | -105.2337
11.0581 | -105.2317
11.0618 | -105.2304
11.0655 | -105.2284
11.0692 | -105.2287
11.0729 | -105.2299
11.0766 | -105.2311
11.0803 | -105.2325
11.084 | -105.2358
11.0876 | -105.2463
11.0913 | -105.2555
11.095 | -105.2649
11.0987 | -105.2745
11.1024 | -105.2854
11.1061 | -105.3001
11.1098 | -105.3167
11.1135 | -105.3345
11.1172 | -105.3529
11.1208 | -105.3706
11.1245 | -105.3938
11.1282 | -105.4171
11.1319 | -105.443
11.1356 | -105.4705
11.1393 | -105.4993
11.143 | -105.5264
11.1467 | -105.5535
11.1504 | -105.5771
11.154 | -105.5982
11.1577 | -105.6188
11.1614 | -105.6403
11.1651 | -105.6632
11.1688 | -105.6831
11.1725 | -105.6959
11.1762 | -105.7123
11.1799 | -105.7271
11.1835 | -105.7417
11.1872 | -105.756
11.1909 | -105.7692
11.1946 | -105.7788
11.1983 | -105.7833
11.202 | -105.7869
11.2057 | -105.7949
11.2094 | -105.8024
11.2131 | -105.8072
11.2167 | -105.8089
11.2204 | -105.8047
11.2241 | -105.8029
11.2278 | -105.8041
11.2315 | -105.8011
11.2352 | -105.7965
11.2389 | -105.7908
11.2426 | -105.7859
11.2463 | -105.7812
11.2499 | -105.7749
11.2536 | -105.7729
11.2573 | -105.7707
11.261 | -105.7642
11.2647 | -105.7545
11.2684 | -105.7477
11.2721 | -105.74
11.2758 | -105.7314
11.2795 | -105.7269
11.2831 | -105.7202
11.2868 | -105.7123
11.2905 | -105.7071
11.2942 | -105.7013
11.2979 | -105.699
11.3016 | -105.6983
11.3053 | -105.6946
11.309 | -105.6917
11.3126 | -105.6877
11.3163 | -105.6895
11.32 | -105.6869
11.3237 | -105.6834
11.3274 | -105.6788
11.3311 | -105.6756
11.3348 | -105.6725
11.3385 | -105.6683
11.3422 | -105.6611
11.3458 | -105.6515
11.3495 | -105.6378
11.3532 | -105.6221
11.3569 | -105.605
11.3606 | -105.5875
11.3643 | -105.5688
11.368 | -105.5463
11.3717 | -105.5209
11.3754 | -105.4995
11.379 | -105.4802
11.3827 | -105.464
11.3864 | -105.4508
11.3901 | -105.4361
11.3938 | -105.4211
11.3975 | -105.4109
11.4012 | -105.4018
11.4049 | -105.3947
11.4085 | -105.3847
11.4122 | -105.377
11.4159 | -105.3728
11.4196 | -105.3694
11.4233 | -105.3687
11.427 | -105.372
11.4307 | -105.3791
11.4344 | -105.384
11.4381 | -105.3889
11.4417 | -105.3989
11.4454 | -105.4085
11.4491 | -105.4208
11.4528 | -105.4314
11.4565 | -105.4386
11.4602 | -105.4467
11.4639 | -105.454
11.4676 | -105.4599
11.4713 | -105.4696
11.4749 | -105.4765
11.4786 | -105.4852
11.4823 | -105.4891
11.486 | -105.4993
11.4897 | -105.5057
11.4934 | -105.511
11.4971 | -105.5197
11.5008 | -105.5276
11.5044 | -105.5361
11.5081 | -105.5456
11.5118 | -105.5571
11.5155 | -105.5678
11.5192 | -105.5772
11.5229 | -105.5835
11.5266 | -105.5941
11.5303 | -105.6028
11.534 | -105.6081
11.5376 | -105.6147
11.5413 | -105.6192
11.545 | -105.6257
11.5487 | -105.6334
11.5524 | -105.642
11.5561 | -105.6493
11.5598 | -105.6561
11.5635 | -105.6594
11.5672 | -105.6649
11.5708 | -105.6714
11.5745 | -105.6825
11.5782 | -105.6925
11.5819 | -105.7006
11.5856 | -105.7082
11.5893 | -105.7149
11.593 | -105.7205
11.5967 | -105.7253
11.6004 | -105.7267
11.604 | -105.7273
11.6077 | -105.7254
11.6114 | -105.7253
11.6151 | -105.7231
11.6188 | -105.7221
11.6225 | -105.722
11.6262 | -105.7153
11.6299 | -105.7066
11.6335 | -105.6998
11.6372 | -105.6872
11.6409 | -105.673
11.6446 | -105.6545
11.6483 | -105.6344
11.652 | -105.6126
11.6557 | -105.5907
11.6594 | -105.5675
11.6631 | -105.5439
11.6667 | -105.5247
11.6704 | -105.5061
11.6741 | -105.4882
11.6778 | -105.471
11.6815 | -105.4574
11.6852 | -105.4447
11.6889 | -105.4348
11.6926 | -105.4253
11.6963 | -105.4198
11.6999 | -105.4165
11.7036 | -105.4195
11.7073 | -105.4148
11.711 | -105.4124
11.7147 | -105.4117
11.7184 | -105.4105
11.7221 | -105.4121
11.7258 | -105.4169
11.7294 | -105.4239
11.7331 | -105.4346
11.7368 | -105.4445
11.7405 | -105.4514
11.7442 | -105.4584
11.7479 | -105.4674
11.7516 | -105.4773
11.7553 | -105.4884
11.759 | -105.5038
11.7626 | -105.5206
11.7663 | -105.5342
11.77 | -105.5524
11.7737 | -105.5713
11.7774 | -105.5863
11.7811 | -105.596
11.7848 | -105.6072
11.7885 | -105.6204
11.7922 | -105.6357
11.7958 | -105.6506
11.7995 | -105.6629
11.8032 | -105.6733
11.8069 | -105.6841
11.8106 | -105.6952
11.8143 | -105.7065
11.818 | -105.7158
11.8217 | -105.7258
11.8253 | -105.7356
11.829 | -105.7448
11.8327 | -105.7512
11.8364 | -105.755
11.8401 | -105.7583
11.8438 | -105.7632
11.8475 | -105.7671
11.8512 | -105.7662
11.8549 | -105.7656
11.8585 | -105.7668
11.8622 | -105.7656
11.8659 | -105.7613
11.8696 | -105.7593
11.8733 | -105.7569
11.877 | -105.7549
11.8807 | -105.7537
11.8844 | -105.7526
11.8881 | -105.7507
11.8917 | -105.7511
11.8954 | -105.7522
11.8991 | -105.7514
11.9028 | -105.7545
11.9065 | -105.7608
11.9102 | -105.7685
11.9139 | -105.7794
11.9176 | -105.7908
11.9213 | -105.8044
11.9249 | -105.8241
11.9286 | -105.8462
11.9323 | -105.8674
11.936 | -105.887
11.9397 | -105.9101
11.9434 | -105.9341
11.9471 | -105.9567
11.9508 | -105.9799
11.9544 | -106.0041
11.9581 | -106.0285
11.9618 | -106.0448
11.9655 | -106.0531
11.9692 | -106.0625
11.9729 | -106.0648
11.9766 | -106.0595
11.9803 | -106.0494
11.984 | -106.035
11.9876 | -106.0183
11.9913 | -105.9963
11.995 | -105.9666
11.9987 | -105.9334
12.0024 | -105.8963
12.0061 | -105.8587
12.0098 | -105.8186
12.0135 | -105.7791
12.0172 | -105.7386
12.0208 | -105.7011
12.0245 | -105.6682
12.0282 | -105.6368
12.0319 | -105.6065
12.0356 | -105.5797
12.0393 | -105.5562
12.043 | -105.5361
12.0467 | -105.519
12.0503 | -105.5046
12.054 | -105.4936
12.0577 | -105.4862
12.0614 | -105.4854
12.0651 | -105.4858
12.0688 | -105.4868
12.0725 | -105.4933
12.0762 | -105.4971
12.0799 | -105.5025
12.0835 | -105.5125
12.0872 | -105.5244
12.0909 | -105.5349
12.0946 | -105.5442
12.0983 | -105.5507
12.102 | -105.5594
12.1057 | -105.5672
12.1094 | -105.5774
12.1131 | -105.5877
12.1167 | -105.5968
12.1204 | -105.6065
12.1241 | -105.619
12.1278 | -105.6279
12.1315 | -105.6351
12.1352 | -105.6477
12.1389 | -105.6624
12.1426 | -105.679
12.1462 | -105.6975
12.1499 | -105.7186
12.1536 | -105.7421
12.1573 | -105.7644
12.161 | -105.7823
12.1647 | -105.8033
12.1684 | -105.825
12.1721 | -105.8475
12.1758 | -105.8703
12.1794 | -105.8935
12.1831 | -105.9179
12.1868 | -105.938
12.1905 | -105.957
12.1942 | -105.9729
12.1979 | -105.9881
12.2016 | -106
12.2053 | -106.0112
12.209 | -106.0207
12.2126 | -106.0284
12.2163 | -106.0311
12.22 | -106.0317
12.2237 | -106.033
12.2274 | -106.0298
12.2311 | -106.0261
12.2348 | -106.018
12.2385 | -106.0056
12.2422 | -105.9972
12.2458 | -105.9903
12.2495 | -105.9824
12.2532 | -105.9739
12.2569 | -105.963
12.2606 | -105.9519
12.2643 | -105.943
12.268 | -105.9381
12.2717 | -105.9316
12.2753 | -105.9266
12.279 | -105.9221
12.2827 | -105.9194
12.2864 | -105.9175
12.2901 | -105.9146
12.2938 | -105.9141
12.2975 | -105.9167
12.3012 | -105.9178
12.3049 | -105.923
12.3085 | -105.9306
12.3122 | -105.937
12.3159 | -105.9463
12.3196 | -105.9564
12.3233 | -105.9682
12.327 | -105.9802
12.3307 | -105.9928
12.3344 | -106.0045
12.3381 | -106.0171
12.3417 | -106.0295
12.3454 | -106.0432
12.3491 | -106.0551
12.3528 | -106.0656
12.3565 | -106.0688
12.3602 | -106.0711
12.3639 | -106.0742
12.3676 | -106.0751
12.3712 | -106.0742
12.3749 | -106.0699
12.3786 | -106.0647
12.3823 | -106.054
12.386 | -106.0429
12.3897 | -106.0307
12.3934 | -106.0148
12.3971 | -105.9985
12.4008 | -105.9821
12.4044 | -105.9687
12.4081 | -105.9518
12.4118 | -105.9305
12.4155 | -105.9089
12.4192 | -105.8915
12.4229 | -105.8727
12.4266 | -105.8548
12.4303 | -105.8403
12.434 | -105.8271
12.4376 | -105.815
12.4413 | -105.8009
12.445 | -105.7882
12.4487 | -105.7775
12.4524 | -105.7684
12.4561 | -105.7601
12.4598 | -105.7536
12.4635 | -105.7491
12.4671 | -105.7459
12.4708 | -105.7448
12.4745 | -105.7426
12.4782 | -105.7398
12.4819 | -105.7391
12.4856 | -105.7373
12.4893 | -105.7367
12.493 | -105.7392
12.4967 | -105.7405
12.5003 | -105.7397
12.504 | -105.7385
12.5077 | -105.7391
12.5114 | -105.7395
12.5151 | -105.7409
12.5188 | -105.7459
12.5225 | -105.7492
12.5262 | -105.7481
12.5299 | -105.747
12.5335 | -105.7472
12.5372 | -105.7483
12.5409 | -105.7468
12.5446 | -105.7443
12.5483 | -105.7411
12.552 | -105.7369
12.5557 | -105.7343
12.5594 | -105.7309
12.5631 | -105.7282
12.5667 | -105.7249
12.5704 | -105.7227
12.5741 | -105.7209
12.5778 | -105.718
12.5815 | -105.7138
12.5852 | -105.7073
12.5889 | -105.7039
12.5926 | -105.707
12.5962 | -105.7121
12.5999 | -105.7166
12.6036 | -105.7231
12.6073 | -105.7294
12.611 | -105.736
12.6147 | -105.747
12.6184 | -105.7601
12.6221 | -105.7764
12.6258 | -105.7937
12.6294 | -105.8096
12.6331 | -105.8273
12.6368 | -105.8465
12.6405 | -105.8636
12.6442 | -105.8792
12.6479 | -105.8974
12.6516 | -105.9126
12.6553 | -105.9299
12.659 | -105.9423
12.6626 | -105.9544
12.6663 | -105.9655
12.67 | -105.9747
12.6737 | -105.9803
12.6774 | -105.9844
12.6811 | -105.9886
12.6848 | -105.9933
12.6885 | -105.9946
12.6921 | -105.9924
12.6958 | -105.9886
12.6995 | -105.9827
12.7032 | -105.976
12.7069 | -105.9712
12.7106 | -105.9648
12.7143 | -105.9553
12.718 | -105.9462
12.7217 | -105.9383
12.7253 | -105.9327
12.729 | -105.9256
12.7327 | -105.9166
12.7364 | -105.9088
12.7401 | -105.9043
12.7438 | -105.8993
12.7475 | -105.8953
12.7512 | -105.892
12.7549 | -105.8883
12.7585 | -105.8822
12.7622 | -105.8769
12.7659 | -105.8766
12.7696 | -105.8685
12.7733 | -105.8641
12.777 | -105.8566
12.7807 | -105.8456
12.7844 | -105.8363
12.788 | -105.8288
12.7917 | -105.8183
12.7954 | -105.8068
12.7991 | -105.7962
12.8028 | -105.7828
12.8065 | -105.7693
12.8102 | -105.7563
12.8139 | -105.742
12.8176 | -105.7253
12.8212 | -105.7041
12.8249 | -105.6854
12.8286 | -105.669
12.8323 | -105.6491
12.836 | -105.6319
12.8397 | -105.6131
12.8434 | -105.5995
12.8471 | -105.5854
12.8508 | -105.5727
12.8544 | -105.5579
12.8581 | -105.5437
12.8618 | -105.5337
12.8655 | -105.5236
12.8692 | -105.5111
12.8729 | -105.5041
12.8766 | -105.5027
12.8803 | -105.4995
12.884 | -105.4965
12.8876 | -105.4958
12.8913 | -105.4942
12.895 | -105.4932
12.8987 | -105.495
12.9024 | -105.4988
12.9061 | -105.5044
12.9098 | -105.5082
12.9135 | -105.5127
12.9171 | -105.5218
12.9208 | -105.5294
12.9245 | -105.5379
12.9282 | -105.5485
12.9319 | -105.5617
12.9356 | -105.5736
12.9393 | -105.5891
12.943 | -105.6094
12.9467 | -105.6295
12.9503 | -105.6471
12.954 | -105.6674
12.9577 | -105.6932
12.9614 | -105.7201
12.9651 | -105.7448
12.9688 | -105.773
12.9725 | -105.7991
12.9762 | -105.8243
12.9799 | -105.8534
12.9835 | -105.8816
12.9872 | -105.9077
12.9909 | -105.9354
12.9946 | -105.9604
12.9983 | -105.9801
13.002 | -106.0009
13.0057 | -106.0206
13.0094 | -106.0362
13.013 | -106.045
13.0167 | -106.0506
13.0204 | -106.0558
13.0241 | -106.0584
13.0278 | -106.0578
13.0315 | -106.0513
13.0352 | -106.0425
13.0389 | -106.0327
13.0426 | -106.0239
13.0462 | -106.014
13.0499 | -106.0001
13.0536 | -105.9846
13.0573 | -105.9727
13.061 | -105.9562
13.0647 | -105.9374
13.0684 | -105.9232
13.0721 | -105.9088
13.0758 | -105.8923
13.0794 | -105.8763
13.0831 | -105.861
13.0868 | -105.8442
13.0905 | -105.8288
13.0942 | -105.815
13.0979 | -105.8026
13.1016 | -105.7925
13.1053 | -105.7809
13.1089 | -105.7725
13.1126 | -105.7648
13.1163 | -105.7593
13.12 | -105.7545
13.1237 | -105.7534
13.1274 | -105.7531
13.1311 | -105.7516
13.1348 | -105.7503
13.1385 | -105.7531
13.1421 | -105.7552
13.1458 | -105.7565
13.1495 | -105.7576
13.1532 | -105.7598
13.1569 | -105.7628
13.1606 | -105.7667
13.1643 | -105.7704
13.168 | -105.7728
13.1717 | -105.7742
13.1753 | -105.778
13.179 | -105.782
13.1827 | -105.7839
13.1864 | -105.7905
13.1901 | -105.7997
13.1938 | -105.8091
13.1975 | -105.8188
13.2012 | -105.8278
13.2049 | -105.8335
13.2085 | -105.8404
13.2122 | -105.8483
13.2159 | -105.8563
13.2196 | -105.864
13.2233 | -105.8713
13.227 | -105.8781
13.2307 | -105.8841
13.2344 | -105.8865
13.238 | -105.8913
13.2417 | -105.8941
13.2454 | -105.8937
13.2491 | -105.8952
13.2528 | -105.8963
13.2565 | -105.898
13.2602 | -105.8989
13.2639 | -105.8999
13.2676 | -105.901
13.2712 | -105.9029
13.2749 | -105.9076
13.2786 | -105.9093
13.2823 | -105.9072
13.286 | -105.9065
13.2897 | -105.9044
13.2934 | -105.8996
13.2971 | -105.8958
13.3008 | -105.8835
13.3044 | -105.8755
13.3081 | -105.8678
13.3118 | -105.8578
13.3155 | -105.8493
13.3192 | -105.84
13.3229 | -105.8268
13.3266 | -105.8137
13.3303 | -105.7944
13.3339 | -105.7747
13.3376 | -105.7561
13.3413 | -105.7364
13.345 | -105.7189
13.3487 | -105.7002
13.3524 | -105.6803
13.3561 | -105.6615
13.3598 | -105.6439
13.3635 | -105.6249
13.3671 | -105.6071
13.3708 | -105.5913
13.3745 | -105.5799
13.3782 | -105.5661
13.3819 | -105.5482
13.3856 | -105.5345
13.3893 | -105.5264
13.393 | -105.5202
13.3967 | -105.5189
13.4003 | -105.5155
13.404 | -105.5116
13.4077 | -105.5087
13.4114 | -105.5116
13.4151 | -105.5158
13.4188 | -105.5196
13.4225 | -105.5255
13.4262 | -105.5359
13.4298 | -105.5432
13.4335 | -105.5527
13.4372 | -105.5659
13.4409 | -105.5767
13.4446 | -105.5889
13.4483 | -105.6019
13.452 | -105.6168
13.4557 | -105.6344
13.4594 | -105.6521
13.463 | -105.6703
13.4667 | -105.6874
13.4704 | -105.7063
13.4741 | -105.7274
13.4778 | -105.7486
13.4815 | -105.7738
13.4852 | -105.7996
13.4889 | -105.8229
13.4926 | -105.8474
13.4962 | -105.8726
13.4999 | -105.8963
13.5036 | -105.9165
13.5073 | -105.9383
13.511 | -105.9594
13.5147 | -105.9807
13.5184 | -106.0036
13.5221 | -106.0243
13.5258 | -106.049
13.5294 | -106.0716
13.5331 | -106.0889
13.5368 | -106.106
13.5405 | -106.1209
13.5442 | -106.1341
13.5479 | -106.1451
13.5516 | -106.1508
13.5553 | -106.1547
13.5589 | -106.1624
13.5626 | -106.1684
13.5663 | -106.1722
13.57 | -106.1736
13.5737 | -106.1764
13.5774 | -106.1783
13.5811 | -106.1744
13.5848 | -106.1693
13.5885 | -106.1641
13.5921 | -106.1561
13.5958 | -106.1482
13.5995 | -106.1421
13.6032 | -106.1385
13.6069 | -106.1301
13.6106 | -106.1213
13.6143 | -106.1124
13.618 | -106.1033
13.6217 | -106.093
13.6253 | -106.0881
13.629 | -106.0847
13.6327 | -106.0797
13.6364 | -106.0767
13.6401 | -106.0727
13.6438 | -106.0694
13.6475 | -106.0688
13.6512 | -106.0675
13.6548 | -106.0697
13.6585 | -106.0726
13.6622 | -106.0765
13.6659 | -106.0807
13.6696 | -106.0863
13.6733 | -106.0921
13.677 | -106.0962
13.6807 | -106.1022
13.6844 | -106.1042
13.688 | -106.1051
13.6917 | -106.1063
13.6954 | -106.103
13.6991 | -106.1015
13.7028 | -106.0948
13.7065 | -106.0857
13.7102 | -106.0745
13.7139 | -106.0579
13.7176 | -106.0413
13.7212 | -106.0236
13.7249 | -106.0009
13.7286 | -105.9767
13.7323 | -105.9523
13.736 | -105.9261
13.7397 | -105.8999
13.7434 | -105.8764
13.7471 | -105.8552
13.7507 | -105.8335
13.7544 | -105.8079
13.7581 | -105.782
13.7618 | -105.7576
13.7655 | -105.7344
13.7692 | -105.7118
13.7729 | -105.6934
13.7766 | -105.6774
13.7803 | -105.6604
13.7839 | -105.6455
13.7876 | -105.6344
13.7913 | -105.6264
13.795 | -105.6196
13.7987 | -105.6127
13.8024 | -105.6094
13.8061 | -105.6094
13.8098 | -105.6131
13.8135 | -105.6192
13.8171 | -105.6297
13.8208 | -105.6371
13.8245 | -105.6477
13.8282 | -105.6572
13.8319 | -105.6723
13.8356 | -105.689
13.8393 | -105.7083
13.843 | -105.7271
13.8467 | -105.7449
13.8503 | -105.7627
13.854 | -105.78
13.8577 | -105.7997
13.8614 | -105.819
13.8651 | -105.8362
13.8688 | -105.8559
13.8725 | -105.8774
13.8762 | -105.8981
13.8798 | -105.9173
13.8835 | -105.9403
13.8872 | -105.9625
13.8909 | -105.982
13.8946 | -106.0047
13.8983 | -106.024
13.902 | -106.0405
13.9057 | -106.0541
13.9094 | -106.0668
13.913 | -106.0825
13.9167 | -106.0957
13.9204 | -106.1064
13.9241 | -106.1201
13.9278 | -106.1319
13.9315 | -106.1389
13.9352 | -106.1444
13.9389 | -106.1483
13.9426 | -106.1501
13.9462 | -106.1485
13.9499 | -106.1478
13.9536 | -106.1441
13.9573 | -106.129
13.961 | -106.0845
13.9647 | -106.0256
13.9684 | -105.9898
13.9721 | -105.9543
13.9757 | -105.9161
13.9794 | -105.8743
13.9831 | -105.8265
13.9868 | -105.7729
13.9905 | -105.7136
13.9942 | -105.652
13.9979 | -105.5872
14.0016 | -105.5183
14.0053 | -105.4464
14.0089 | -105.3719
14.0126 | -105.2941
14.0163 | -105.2156
14.02 | -105.1392
14.0237 | -105.064
14.0274 | -104.9922
14.0311 | -104.9241
14.0348 | -104.859
14.0385 | -104.7987
14.0421 | -104.7435
14.0458 | -104.6954
14.0495 | -104.6551
14.0532 | -104.625
14.0569 | -104.6042
14.0606 | -104.5903
14.0643 | -104.5842
14.068 | -104.5877
14.0716 | -104.5961
14.0753 | -104.6113
14.079 | -104.6359
14.0827 | -104.6672
14.0864 | -104.7039
14.0901 | -104.7479
14.0938 | -104.7978
14.0975 | -104.8493
14.1012 | -104.9064
14.1048 | -104.9663
14.1085 | -105.0265
14.1122 | -105.0852
14.1159 | -105.1419
14.1196 | -105.1971
14.1233 | -105.2487
14.127 | -105.2965
14.1307 | -105.3412
14.1344 | -105.3831
14.138 | -105.4213
14.1417 | -105.4544
14.1454 | -105.4852
14.1491 | -105.5143
14.1528 | -105.5354
14.1565 | -105.5597
14.1602 | -105.5868
14.1639 | -105.6299
14.1676 | -105.6694
14.1712 | -105.6756
14.1749 | -105.6766
14.1786 | -105.6768
14.1823 | -105.6789
14.186 | -105.6798
14.1897 | -105.6813
14.1934 | -105.6859
14.1971 | -105.689
14.2007 | -105.6921
14.2044 | -105.6968
14.2081 | -105.7031
14.2118 | -105.7101
14.2155 | -105.7195
14.2192 | -105.7307
14.2229 | -105.7429
14.2266 | -105.7545
14.2303 | -105.769
14.2339 | -105.7809
14.2376 | -105.7939
14.2413 | -105.8097
14.245 | -105.8273
14.2487 | -105.8461
14.2524 | -105.8653
14.2561 | -105.8814
14.2598 | -105.8988
14.2635 | -105.9171
14.2671 | -105.9393
14.2708 | -105.9626
14.2745 | -105.9842
14.2782 | -106.003
14.2819 | -106.0182
14.2856 | -106.0343
14.2893 | -106.0443
14.293 | -106.0506
14.2966 | -106.0559
14.3003 | -106.062
14.304 | -106.0617
14.3077 | -106.0613
14.3114 | -106.0609
14.3151 | -106.0613
14.3188 | -106.0612
14.3225 | -106.0578
14.3262 | -106.0528
14.3298 | -106.0478
14.3335 | -106.0418
14.3372 | -106.0361
14.3409 | -106.0301
14.3446 | -106.0233
14.3483 | -106.0159
14.352 | -106.0101
14.3557 | -106.0039
14.3594 | -105.9945
14.363 | -105.9877
14.3667 | -105.9846
14.3704 | -105.9778
14.3741 | -105.9746
14.3778 | -105.9728
14.3815 | -105.9699
14.3852 | -105.9667
14.3889 | -105.963
14.3925 | -105.9594
14.3962 | -105.9549
14.3999 | -105.9509
14.4036 | -105.9473
14.4073 | -105.944
14.411 | -105.9413
14.4147 | -105.9391
14.4184 | -105.9359
14.4221 | -105.9312
14.4257 | -105.9264
14.4294 | -105.9238
14.4331 | -105.9204
14.4368 | -105.9166
14.4405 | -105.915
14.4442 | -105.9099
14.4479 | -105.9052
14.4516 | -105.9051
14.4553 | -105.9051
14.4589 | -105.9037
14.4626 | -105.905
14.4663 | -105.9036
14.47 | -105.9036
14.4737 | -105.9054
14.4774 | -105.9058
14.4811 | -105.9079
14.4848 | -105.9102
14.4885 | -105.915
14.4921 | -105.92
14.4958 | -105.929
14.4995 | -105.938
14.5032 | -105.9456
14.5069 | -105.9541
14.5106 | -105.9641
14.5143 | -105.9718
14.518 | -105.9796
14.5216 | -105.9864
14.5253 | -105.9938
14.529 | -106.002
14.5327 | -106.0117
14.5364 | -106.0216
14.5401 | -106.0299
14.5438 | -106.0388
14.5475 | -106.0461
14.5512 | -106.0549
14.5548 | -106.0585
14.5585 | -106.0613
14.5622 | -106.0664
14.5659 | -106.0679
14.5696 | -106.0665
14.5733 | -106.0653
14.577 | -106.067
14.5807 | -106.0634
14.5844 | -106.0553
14.588 | -106.0491
14.5917 | -106.0434
14.5954 | -106.0358
14.5991 | -106.0273
14.6028 | -106.0225
14.6065 | -106.0159
14.6102 | -106.0067
14.6139 | -105.9965
14.6175 | -105.9885
14.6212 | -105.9795
14.6249 | -105.9689
14.6286 | -105.9588
14.6323 | -105.9521
14.636 | -105.9463
14.6397 | -105.9361
14.6434 | -105.9269
14.6471 | -105.9225
14.6507 | -105.9207
14.6544 | -105.9199
14.6581 | -105.917
14.6618 | -105.9133
14.6655 | -105.9115
14.6692 | -105.91
14.6729 | -105.908
14.6766 | -105.9073
14.6803 | -105.9057
14.6839 | -105.9025
14.6876 | -105.9042
14.6913 | -105.9067
14.695 | -105.9105
14.6987 | -105.911
14.7024 | -105.9104
14.7061 | -105.9111
14.7098 | -105.9113
14.7134 | -105.9139
14.7171 | -105.9141
14.7208 | -105.9131
14.7245 | -105.9152
14.7282 | -105.9157
14.7319 | -105.9163
14.7356 | -105.9185
14.7393 | -105.9218
14.743 | -105.9233
14.7466 | -105.9236
14.7503 | -105.9243
14.754 | -105.9235
14.7577 | -105.9254
14.7614 | -105.9261
14.7651 | -105.9285
14.7688 | -105.9319
14.7725 | -105.9349
14.7762 | -105.9405
14.7798 | -105.9447
14.7835 | -105.9452
14.7872 | -105.9451
14.7909 | -105.9484
14.7946 | -105.955
14.7983 | -105.9589
14.802 | -105.9598
14.8057 | -105.9629
14.8094 | -105.9645
14.813 | -105.9644
14.8167 | -105.964
14.8204 | -105.958
14.8241 | -105.9495
14.8278 | -105.9387
14.8315 | -105.9287
14.8352 | -105.9208
14.8389 | -105.9098
14.8425 | -105.8956
14.8462 | -105.8798
14.8499 | -105.8688
14.8536 | -105.8568
14.8573 | -105.8384
14.861 | -105.821
14.8647 | -105.804
14.8684 | -105.7887
14.8721 | -105.7727
14.8757 | -105.7585
14.8794 | -105.7477
14.8831 | -105.7366
14.8868 | -105.7273
14.8905 | -105.7203
14.8942 | -105.7136
14.8979 | -105.716
14.9016 | -105.7096
14.9053 | -105.7062
14.9089 | -105.7039
14.9126 | -105.7037
14.9163 | -105.7059
14.92 | -105.7112
14.9237 | -105.7168
14.9274 | -105.7239
14.9311 | -105.7322
14.9348 | -105.7439
14.9384 | -105.7559
14.9421 | -105.7682
14.9458 | -105.7805
14.9495 | -105.7889
14.9532 | -105.7968
14.9569 | -105.8092
14.9606 | -105.8216
14.9643 | -105.8337
14.968 | -105.8441
14.9716 | -105.8498
14.9753 | -105.8531
14.979 | -105.8582
14.9827 | -105.8586
14.9864 | -105.8539
14.9901 | -105.8473
14.9938 | -105.843
14.9975 | -105.8367
15.0012 | -105.8265
15.0048 | -105.8127
15.0085 | -105.7955
15.0122 | -105.7782
15.0159 | -105.7625
15.0196 | -105.7453
15.0233 | -105.7281
15.027 | -105.7113
15.0307 | -105.6945
15.0343 | -105.6824
15.038 | -105.6711
15.0417 | -105.6582
15.0454 | -105.643
15.0491 | -105.6267
15.0528 | -105.6142
15.0565 | -105.6025
15.0602 | -105.5917
15.0639 | -105.5804
15.0675 | -105.5729
15.0712 | -105.5642
15.0749 | -105.5538
15.0786 | -105.5449
15.0823 | -105.5371
15.086 | -105.531
15.0897 | -105.5273
15.0934 | -105.5265
15.0971 | -105.529
15.1007 | -105.5286
15.1044 | -105.5294
15.1081 | -105.5324
15.1118 | -105.5359
15.1155 | -105.543
15.1192 | -105.5502
15.1229 | -105.554
15.1266 | -105.5589
15.1303 | -105.5645
15.1339 | -105.5709
15.1376 | -105.5746
15.1413 | -105.5788
15.145 | -105.5816
15.1487 | -105.5846
15.1524 | -105.5876
15.1561 | -105.5917
15.1598 | -105.5957
15.1634 | -105.5991
15.1671 | -105.6044
15.1708 | -105.6078
15.1745 | -105.6082
15.1782 | -105.6091
15.1819 | -105.6114
15.1856 | -105.6162
15.1893 | -105.6192
15.193 | -105.6233
15.1966 | -105.6279
15.2003 | -105.6348
15.204 | -105.6444
15.2077 | -105.6527
15.2114 | -105.658
15.2151 | -105.6675
15.2188 | -105.676
15.2225 | -105.6849
15.2262 | -105.6935
15.2298 | -105.7019
15.2335 | -105.7096
15.2372 | -105.7155
15.2409 | -105.7206
15.2446 | -105.7236
15.2483 | -105.729
15.252 | -105.732
15.2557 | -105.7319
15.2593 | -105.732
15.263 | -105.7331
15.2667 | -105.7345
15.2704 | -105.7347
15.2741 | -105.7346
15.2778 | -105.7327
15.2815 | -105.7288
15.2852 | -105.7219
15.2889 | -105.7181
15.2925 | -105.7139
15.2962 | -105.7064
15.2999 | -105.6971
15.3036 | -105.6895
15.3073 | -105.6798
15.311 | -105.6654
15.3147 | -105.651
15.3184 | -105.6419
15.3221 | -105.6293
15.3257 | -105.6142
15.3294 | -105.6024
15.3331 | -105.5893
15.3368 | -105.5743
15.3405 | -105.5642
15.3442 | -105.5524
15.3479 | -105.541
15.3516 | -105.5316
15.3552 | -105.5241
15.3589 | -105.5168
15.3626 | -105.5099
15.3663 | -105.5048
15.37 | -105.4988
15.3737 | -105.4936
15.3774 | -105.4908
15.3811 | -105.4893
15.3848 | -105.4897
15.3884 | -105.4895
15.3921 | -105.4914
15.3958 | -105.4949
15.3995 | -105.4978
15.4032 | -105.4999
15.4069 | -105.5054
15.4106 | -105.5094
15.4143 | -105.5114
15.418 | -105.5137
15.4216 | -105.5184
15.4253 | -105.5217
15.429 | -105.5161
15.4327 | -105.5161
15.4364 | -105.5126
15.4401 | -105.51
15.4438 | -105.5095
15.4475 | -105.5066
15.4512 | -105.4993
15.4548 | -105.494
15.4585 | -105.488
15.4622 | -105.4786
15.4659 | -105.4698
15.4696 | -105.4575
15.4733 | -105.4469
15.477 | -105.4349
15.4807 | -105.4218
15.4843 | -105.4039
15.488 | -105.3869
15.4917 | -105.3712
15.4954 | -105.3525
15.4991 | -105.3346
15.5028 | -105.3193
15.5065 | -105.305
15.5102 | -105.2885
15.5139 | -105.2734
15.5175 | -105.2614
15.5212 | -105.2501
15.5249 | -105.2356
15.5286 | -105.2219
15.5323 | -105.2111
15.536 | -105.2019
15.5397 | -105.1937
15.5434 | -105.1915
15.5471 | -105.1915
15.5507 | -105.1876
15.5544 | -105.1851
15.5581 | -105.1871
15.5618 | -105.1904
15.5655 | -105.1943
15.5692 | -105.197
15.5729 | -105.1988
15.5766 | -105.2016
15.5802 | -105.2058
15.5839 | -105.2098
15.5876 | -105.2133
15.5913 | -105.2191
15.595 | -105.2234
15.5987 | -105.2245
15.6024 | -105.2274
15.6061 | -105.2275
15.6098 | -105.2273
15.6134 | -105.2257
15.6171 | -105.2222
15.6208 | -105.2174
15.6245 | -105.211
15.6282 | -105.2039
15.6319 | -105.1908
15.6356 | -105.1776
15.6393 | -105.1668
15.643 | -105.1545
15.6466 | -105.1423
15.6503 | -105.1316
15.654 | -105.1178
15.6577 | -105.1031
15.6614 | -105.0888
15.6651 | -105.0764
15.6688 | -105.0644
15.6725 | -105.0535
15.6761 | -105.0433
15.6798 | -105.0323
15.6835 | -105.0239
15.6872 | -105.0172
15.6909 | -105.0102
15.6946 | -105.0075
15.6983 | -105.0036
15.702 | -105.002
15.7057 | -104.9995
15.7093 | -105.0009
15.713 | -105.0006
15.7167 | -105.0001
15.7204 | -105.0029
15.7241 | -105.0074
15.7278 | -105.0137
15.7315 | -105.0202
15.7352 | -105.0302
15.7389 | -105.04
15.7425 | -105.0501
15.7462 | -105.06
15.7499 | -105.0722
15.7536 | -105.0871
15.7573 | -105.1035
15.761 | -105.1192
15.7647 | -105.1345
15.7684 | -105.1484
15.7721 | -105.1657
15.7757 | -105.1867
15.7794 | -105.2071
15.7831 | -105.2266
15.7868 | -105.2457
15.7905 | -105.2683
15.7942 | -105.2924
15.7979 | -105.313
15.8016 | -105.3328
15.8052 | -105.3513
15.8089 | -105.3692
15.8126 | -105.3869
15.8163 | -105.4002
15.82 | -105.413
15.8237 | -105.4265
15.8274 | -105.4359
15.8311 | -105.4459
15.8348 | -105.455
15.8384 | -105.4575
15.8421 | -105.4606
15.8458 | -105.4638
15.8495 | -105.4644
15.8532 | -105.4621
15.8569 | -105.4647
15.8606 | -105.463
15.8643 | -105.4551
15.868 | -105.4429
15.8716 | -105.4319
15.8753 | -105.4235
15.879 | -105.4141
15.8827 | -105.4052
15.8864 | -105.3946
15.8901 | -105.3811
15.8938 | -105.3659
15.8975 | -105.3501
15.9011 | -105.3383
15.9048 | -105.3275
15.9085 | -105.3166
15.9122 | -105.3085
15.9159 | -105.2991
15.9196 | -105.2886
15.9233 | -105.2837
15.927 | -105.2808
15.9307 | -105.2767
15.9343 | -105.2729
15.938 | -105.2703
15.9417 | -105.2697
15.9454 | -105.2699
15.9491 | -105.2696
15.9528 | -105.2678
15.9565 | -105.2671
15.9602 | -105.2653
15.9639 | -105.2648
15.9675 | -105.2651
15.9712 | -105.2635
15.9749 | -105.2629
15.9786 | -105.2636
15.9823 | -105.2625
15.986 | -105.2619
15.9897 | -105.2613
15.9934 | -105.2622
15.997 | -105.2634
16.0007 | -105.2618
16.0044 | -105.2595
16.0081 | -105.2586
16.0118 | -105.2567
16.0155 | -105.2548
16.0192 | -105.2517
16.0229 | -105.2484
16.0266 | -105.2455
16.0302 | -105.2439
16.0339 | -105.2454
16.0376 | -105.2438
16.0413 | -105.2415
16.045 | -105.239
16.0487 | -105.2344
16.0524 | -105.2307
16.0561 | -105.2297
16.0598 | -105.229
16.0634 | -105.2286
16.0671 | -105.2315
16.0708 | -105.2329
16.0745 | -105.2348
16.0782 | -105.2386
16.0819 | -105.243
16.0856 | -105.2446
16.0893 | -105.2462
16.0929 | -105.2461
16.0966 | -105.2474
16.1003 | -105.2507
16.104 | -105.2533
16.1077 | -105.2563
16.1114 | -105.2571
16.1151 | -105.2589
16.1188 | -105.261
16.1225 | -105.2614
16.1261 | -105.265
16.1298 | -105.2657
16.1335 | -105.2654
16.1372 | -105.2641
16.1409 | -105.2628
16.1446 | -105.2639
16.1483 | -105.2624
16.152 | -105.2592
16.1557 | -105.2571
16.1593 | -105.2597
16.163 | -105.2617
16.1667 | -105.2635
16.1704 | -105.2645
16.1741 | -105.2663
16.1778 | -105.2672
16.1815 | -105.2706
16.1852 | -105.2753
16.1889 | -105.2821
16.1925 | -105.2894
16.1962 | -105.2977
16.1999 | -105.3057
16.2036 | -105.315
16.2073 | -105.3261
16.211 | -105.3345
16.2147 | -105.342
16.2184 | -105.3506
16.222 | -105.3597
16.2257 | -105.3678
16.2294 | -105.3762
16.2331 | -105.3879
16.2368 | -105.3996
16.2405 | -105.4105
16.2442 | -105.421
16.2479 | -105.4311
16.2516 | -105.441
16.2552 | -105.4489
16.2589 | -105.4603
16.2626 | -105.4698
16.2663 | -105.4763
16.27 | -105.4859
16.2737 | -105.494
16.2774 | -105.5028
16.2811 | -105.513
16.2848 | -105.5211
16.2884 | -105.5257
16.2921 | -105.5268
16.2958 | -105.5296
16.2995 | -105.5317
16.3032 | -105.5341
16.3069 | -105.5365
16.3106 | -105.5385
16.3143 | -105.5409
16.3179 | -105.5417
16.3216 | -105.5427
16.3253 | -105.5456
16.329 | -105.5485
16.3327 | -105.5447
16.3364 | -105.5386
16.3401 | -105.5336
16.3438 | -105.5306
16.3475 | -105.5303
16.3511 | -105.5293
16.3548 | -105.5291
16.3585 | -105.5262
16.3622 | -105.5249
16.3659 | -105.5221
16.3696 | -105.5222
16.3733 | -105.5211
16.377 | -105.5191
16.3807 | -105.5176
16.3843 | -105.5148
16.388 | -105.5132
16.3917 | -105.5103
16.3954 | -105.5076
16.3991 | -105.5045
16.4028 | -105.5014
16.4065 | -105.4972
16.4102 | -105.4974
16.4139 | -105.4949
16.4175 | -105.4934
16.4212 | -105.4907
16.4249 | -105.4868
16.4286 | -105.4824
16.4323 | -105.4786
16.436 | -105.475
16.4397 | -105.4715
16.4434 | -105.4689
16.447 | -105.4676
16.4507 | -105.4662
16.4544 | -105.465
16.4581 | -105.4631
16.4618 | -105.4615
16.4655 | -105.4569
16.4692 | -105.4524
16.4729 | -105.4498
16.4766 | -105.4503
16.4802 | -105.4479
16.4839 | -105.4477
16.4876 | -105.4504
16.4913 | -105.4509
16.495 | -105.4522
16.4987 | -105.4476
16.5024 | -105.4473
16.5061 | -105.4476
16.5098 | -105.4479
16.5134 | -105.4464
16.5171 | -105.4458
16.5208 | -105.445
16.5245 | -105.4443
16.5282 | -105.4469
16.5319 | -105.4497
16.5356 | -105.4501
16.5393 | -105.4491
16.5429 | -105.4451
16.5466 | -105.441
16.5503 | -105.4404
16.554 | -105.4432
16.5577 | -105.4432
16.5614 | -105.4397
16.5651 | -105.439
16.5688 | -105.4395
16.5725 | -105.4402
16.5761 | -105.4385
16.5798 | -105.4345
16.5835 | -105.4313
16.5872 | -105.4288
16.5909 | -105.4261
16.5946 | -105.4239
16.5983 | -105.4227
16.602 | -105.425
16.6057 | -105.4252
16.6093 | -105.4264
16.613 | -105.4287
16.6167 | -105.431
16.6204 | -105.435
16.6241 | -105.4398
16.6278 | -105.4448
16.6315 | -105.449
16.6352 | -105.4532
16.6388 | -105.4585
16.6425 | -105.4644
16.6462 | -105.4683
16.6499 | -105.4757
16.6536 | -105.4812
16.6573 | -105.4871
16.661 | -105.4939
16.6647 | -105.5007
16.6684 | -105.5077
16.672 | -105.5147
16.6757 | -105.5223
16.6794 | -105.5303
16.6831 | -105.5369
16.6868 | -105.5406
16.6905 | -105.5448
16.6942 | -105.5487
16.6979 | -105.5508
16.7016 | -105.5524
16.7052 | -105.5522
16.7089 | -105.5519
16.7126 | -105.5461
16.7163 | -105.5425
16.72 | -105.5387
16.7237 | -105.5342
16.7274 | -105.5264
16.7311 | -105.5173
16.7347 | -105.5076
16.7384 | -105.4962
16.7421 | -105.484
16.7458 | -105.4699
16.7495 | -105.4526
16.7532 | -105.4354
16.7569 | -105.4187
16.7606 | -105.3997
16.7643 | -105.3822
16.7679 | -105.3634
16.7716 | -105.3428
16.7753 | -105.3239
16.779 | -105.3071
16.7827 | -105.2933
16.7864 | -105.281
16.7901 | -105.2644
16.7938 | -105.2475
16.7975 | -105.2329
16.8011 | -105.2215
16.8048 | -105.2082
16.8085 | -105.1966
16.8122 | -105.1871
16.8159 | -105.181
16.8196 | -105.1762
16.8233 | -105.167
16.827 | -105.1622
16.8307 | -105.1599
16.8343 | -105.1597
16.838 | -105.1614
16.8417 | -105.1614
16.8454 | -105.163
16.8491 | -105.1652
16.8528 | -105.1681
16.8565 | -105.1738
16.8602 | -105.1777
16.8638 | -105.1807
16.8675 | -105.1843
16.8712 | -105.1913
16.8749 | -105.1998
16.8786 | -105.2059
16.8823 | -105.2091
16.886 | -105.2122
16.8897 | -105.2146
16.8934 | -105.2172
16.897 | -105.221
16.9007 | -105.2248
16.9044 | -105.2264
16.9081 | -105.2277
16.9118 | -105.2286
16.9155 | -105.2291
16.9192 | -105.2272
16.9229 | -105.2236
16.9266 | -105.2195
16.9302 | -105.2157
16.9339 | -105.2117
16.9376 | -105.2085
16.9413 | -105.205
16.945 | -105.2013
16.9487 | -105.1951
16.9524 | -105.1895
16.9561 | -105.1837
16.9597 | -105.1774
16.9634 | -105.1708
16.9671 | -105.1659
16.9708 | -105.162
16.9745 | -105.1579
16.9782 | -105.1537
16.9819 | -105.1456
16.9856 | -105.1399
16.9893 | -105.1373
16.9929 | -105.1334
16.9966 | -105.13
17.0003 | -105.1285
17.004 | -105.1229
17.0077 | -105.1178
17.0114 | -105.1151
17.0151 | -105.1114
17.0188 | -105.1062
17.0225 | -105.103
17.0261 | -105.1037
17.0298 | -105.1009
17.0335 | -105.0999
17.0372 | -105.0973
17.0409 | -105.0945
17.0446 | -105.0894
17.0483 | -105.0831
17.052 | -105.0788
17.0556 | -105.0751
17.0593 | -105.0706
17.063 | -105.0665
17.0667 | -105.0656
17.0704 | -105.0615
17.0741 | -105.0551
17.0778 | -105.049
17.0815 | -105.0409
17.0852 | -105.0334
17.0888 | -105.0277
17.0925 | -105.0226
17.0962 | -105.0195
17.0999 | -105.0157
17.1036 | -105.0138
17.1073 | -105.0124
17.111 | -105.01
17.1147 | -105.0068
17.1184 | -105.0044
17.122 | -105.0028
17.1257 | -105.0032
17.1294 | -105.0054
17.1331 | -105.007
17.1368 | -105.0111
17.1405 | -105.0163
17.1442 | -105.0184
17.1479 | -105.0202
17.1516 | -105.0239
17.1552 | -105.0292
17.1589 | -105.0331
17.1626 | -105.0402
17.1663 | -105.0493
17.17 | -105.0603
17.1737 | -105.0698
17.1774 | -105.0785
17.1811 | -105.0876
17.1847 | -105.0971
17.1884 | -105.105
17.1921 | -105.1128
17.1958 | -105.1223
17.1995 | -105.1313
17.2032 | -105.1377
17.2069 | -105.1439
17.2106 | -105.1493
17.2143 | -105.1547
17.2179 | -105.1627
17.2216 | -105.1706
17.2253 | -105.1768
17.229 | -105.1798
17.2327 | -105.1832
17.2364 | -105.1878
17.2401 | -105.1874
17.2438 | -105.1866
17.2475 | -105.1894
17.2511 | -105.1896
17.2548 | -105.187
17.2585 | -105.1866
17.2622 | -105.183
17.2659 | -105.1782
17.2696 | -105.1713
17.2733 | -105.1634
17.277 | -105.1559
17.2806 | -105.1475
17.2843 | -105.1389
17.288 | -105.1322
17.2917 | -105.1259
17.2954 | -105.117
17.2991 | -105.1036
17.3028 | -105.088
17.3065 | -105.0708
17.3102 | -105.0531
17.3138 | -105.032
17.3175 | -105.0115
17.3212 | -104.9924
17.3249 | -104.9732
17.3286 | -104.9543
17.3323 | -104.9351
17.336 | -104.9155
17.3397 | -104.8964
17.3434 | -104.8708
17.347 | -104.8473
17.3507 | -104.8236
17.3544 | -104.8008
17.3581 | -104.7821
17.3618 | -104.7626
17.3655 | -104.7422
17.3692 | -104.7211
17.3729 | -104.7013
17.3766 | -104.6822
17.3802 | -104.665
17.3839 | -104.6469
17.3876 | -104.6282
17.3913 | -104.6131
17.395 | -104.6018
17.3987 | -104.5899
17.4024 | -104.5786
17.4061 | -104.569
17.4097 | -104.5597
17.4134 | -104.5506
17.4171 | -104.5408
17.4208 | -104.5341
17.4245 | -104.5299
17.4282 | -104.5258
17.4319 | -104.5249
17.4356 | -104.5246
17.4393 | -104.5244
17.4429 | -104.5256
17.4466 | -104.5254
17.4503 | -104.5274
17.454 | -104.5322
17.4577 | -104.5338
17.4614 | -104.5317
17.4651 | -104.533
17.4688 | -104.5354
17.4725 | -104.5365
17.4761 | -104.5382
17.4798 | -104.5407
17.4835 | -104.5423
17.4872 | -104.5438
17.4909 | -104.5452
17.4946 | -104.5451
17.4983 | -104.5454
17.502 | -104.5458
17.5056 | -104.5439
17.5093 | -104.545
17.513 | -104.5452
17.5167 | -104.5455
17.5204 | -104.5445
17.5241 | -104.5414
17.5278 | -104.5403
17.5315 | -104.5404
17.5352 | -104.5378
17.5388 | -104.534
17.5425 | -104.5315
17.5462 | -104.5315
17.5499 | -104.5305
17.5536 | -104.5288
17.5573 | -104.5255
17.561 | -104.5227
17.5647 | -104.5212
17.5684 | -104.5167
17.572 | -104.5139
17.5757 | -104.51
17.5794 | -104.5067
17.5831 | -104.5039
17.5868 | -104.5015
17.5905 | -104.4987
17.5942 | -104.4973
17.5979 | -104.495
17.6015 | -104.4933
17.6052 | -104.4912
17.6089 | -104.4872
17.6126 | -104.4847
17.6163 | -104.4827
17.62 | -104.4815
17.6237 | -104.4793
17.6274 | -104.4758
17.6311 | -104.4729
17.6347 | -104.4701
17.6384 | -104.4663
17.6421 | -104.4598
17.6458 | -104.4514
17.6495 | -104.4423
17.6532 | -104.4326
17.6569 | -104.4233
17.6606 | -104.4138
17.6643 | -104.4025
17.6679 | -104.3891
17.6716 | -104.3758
17.6753 | -104.3632
17.679 | -104.3491
17.6827 | -104.3336
17.6864 | -104.3188
17.6901 | -104.3036
17.6938 | -104.2887
17.6974 | -104.2741
17.7011 | -104.2588
17.7048 | -104.2438
17.7085 | -104.232
17.7122 | -104.2221
17.7159 | -104.2088
17.7196 | -104.1955
17.7233 | -104.1825
17.727 | -104.1688
17.7306 | -104.1547
17.7343 | -104.1408
17.738 | -104.1285
17.7417 | -104.1172
17.7454 | -104.1064
17.7491 | -104.0963
17.7528 | -104.0862
17.7565 | -104.0769
17.7602 | -104.0667
17.7638 | -104.0557
17.7675 | -104.0441
17.7712 | -104.0324
17.7749 | -104.023
17.7786 | -104.0137
17.7823 | -104.0028
17.786 | -103.9943
17.7897 | -103.9892
17.7934 | -103.9813
17.797 | -103.972
17.8007 | -103.9652
17.8044 | -103.9603
17.8081 | -103.9562
17.8118 | -103.9515
17.8155 | -103.9476
17.8192 | -103.9461
17.8229 | -103.9464
17.8265 | -103.9436
17.8302 | -103.9409
17.8339 | -103.938
17.8376 | -103.9367
17.8413 | -103.9363
17.845 | -103.9345
17.8487 | -103.932
17.8524 | -103.9307
17.8561 | -103.9291
17.8597 | -103.9264
17.8634 | -103.9262
17.8671 | -103.9268
17.8708 | -103.9247
17.8745 | -103.9225
17.8782 | -103.9192
17.8819 | -103.9139
17.8856 | -103.9094
17.8893 | -103.9062
17.8929 | -103.9029
17.8966 | -103.8997
17.9003 | -103.8921
17.904 | -103.8864
17.9077 | -103.8793
17.9114 | -103.8705
17.9151 | -103.8604
17.9188 | -103.85
17.9224 | -103.8368
17.9261 | -103.8214
17.9298 | -103.8071
17.9335 | -103.7928
17.9372 | -103.7775
17.9409 | -103.7607
17.9446 | -103.7447
17.9483 | -103.7274
17.952 | -103.709
17.9556 | -103.6908
17.9593 | -103.6713
17.963 | -103.652
17.9667 | -103.6323
17.9704 | -103.6134
17.9741 | -103.5913
17.9778 | -103.5712
17.9815 | -103.5537
17.9852 | -103.534
17.9888 | -103.5143
17.9925 | -103.4964
17.9962 | -103.4775
17.9999 | -103.459
18.0036 | -103.4433
18.0073 | -103.4288
18.011 | -103.4155
18.0147 | -103.403
18.0184 | -103.3917
18.022 | -103.383
18.0257 | -103.3735
18.0294 | -103.3658
18.0331 | -103.3583
18.0368 | -103.3521
18.0405 | -103.3462
18.0442 | -103.341
18.0479 | -103.34
18.0515 | -103.3423
18.0552 | -103.3437
18.0589 | -103.3426
18.0626 | -103.3419
18.0663 | -103.3439
18.07 | -103.3455
18.0737 | -103.3456
18.0774 | -103.345
18.0811 | -103.3472
18.0847 | -103.3511
18.0884 | -103.3537
18.0921 | -103.3594
18.0958 | -103.3606
18.0995 | -103.359
18.1032 | -103.3611
18.1069 | -103.3654
18.1106 | -103.3675
18.1143 | -103.3675
18.1179 | -103.366
18.1216 | -103.3647
18.1253 | -103.3623
18.129 | -103.3605
18.1327 | -103.3593
18.1364 | -103.3574
18.1401 | -103.3551
18.1438 | -103.348
18.1474 | -103.3389
18.1511 | -103.3313
18.1548 | -103.3246
18.1585 | -103.3174
18.1622 | -103.3058
18.1659 | -103.2938
18.1696 | -103.2808
18.1733 | -103.2688
18.177 | -103.2558
18.1806 | -103.2459
18.1843 | -103.2353
18.188 | -103.2192
18.1917 | -103.2035
18.1954 | -103.1878
18.1991 | -103.1724
18.2028 | -103.1546
18.2065 | -103.1359
18.2102 | -103.1189
18.2138 | -103.1024
18.2175 | -103.0858
18.2212 | -103.0679
18.2249 | -103.0515
18.2286 | -103.0316
18.2323 | -103.0161
18.236 | -102.9997
18.2397 | -102.9807
18.2433 | -102.9624
18.247 | -102.9444
18.2507 | -102.9253
18.2544 | -102.9084
18.2581 | -102.8939
18.2618 | -102.8773
18.2655 | -102.8612
18.2692 | -102.8429
18.2729 | -102.825
18.2765 | -102.8091
18.2802 | -102.7955
18.2839 | -102.7847
18.2876 | -102.7704
18.2913 | -102.7557
18.295 | -102.7429
18.2987 | -102.73
18.3024 | -102.7151
18.3061 | -102.7017
18.3097 | -102.6873
18.3134 | -102.6744
18.3171 | -102.6647
18.3208 | -102.6558
18.3245 | -102.6452
18.3282 | -102.6346
18.3319 | -102.6236
18.3356 | -102.6133
18.3392 | -102.6031
18.3429 | -102.5931
18.3466 | -102.5842
18.3503 | -102.574
18.354 | -102.564
18.3577 | -102.5577
18.3614 | -102.5512
18.3651 | -102.5408
18.3688 | -102.5281
18.3724 | -102.5139
18.3761 | -102.5028
18.3798 | -102.495
18.3835 | -102.4863
18.3872 | -102.4759
18.3909 | -102.4648
18.3946 | -102.4519
18.3983 | -102.4408
18.402 | -102.4286
18.4056 | -102.4184
18.4093 | -102.4099
18.413 | -102.4007
18.4167 | -102.3906
18.4204 | -102.3826
18.4241 | -102.374
18.4278 | -102.3647
18.4315 | -102.3554
18.4352 | -102.3461
18.4388 | -102.3382
18.4425 | -102.3249
18.4462 | -102.3117
18.4499 | -102.3031
18.4536 | -102.2935
18.4573 | -102.2849
18.461 | -102.2768
18.4647 | -102.2707
18.4683 | -102.2653
18.472 | -102.2588
18.4757 | -102.2504
18.4794 | -102.2407
18.4831 | -102.2339
18.4868 | -102.2268
18.4905 | -102.2208
18.4942 | -102.2146
18.4979 | -102.2108
18.5015 | -102.209
18.5052 | -102.2075
18.5089 | -102.2063
18.5126 | -102.2065
18.5163 | -102.2038
18.52 | -102.2011
18.5237 | -102.2024
18.5274 | -102.2043
18.5311 | -102.2084
18.5347 | -102.2121
18.5384 | -102.2152
18.5421 | -102.2166
18.5458 | -102.2189
18.5495 | -102.2216
18.5532 | -102.2251
18.5569 | -102.2282
18.5606 | -102.2302
18.5642 | -102.2338
18.5679 | -102.2379
18.5716 | -102.2425
18.5753 | -102.2473
18.579 | -102.2516
18.5827 | -102.2548
18.5864 | -102.2577
18.5901 | -102.2632
18.5938 | -102.2675
18.5974 | -102.2718
18.6011 | -102.2753
18.6048 | -102.2787
18.6085 | -102.2825
18.6122 | -102.2864
18.6159 | -102.2892
18.6196 | -102.2928
18.6233 | -102.3004
18.627 | -102.3052
18.6306 | -102.3124
18.6343 | -102.3204
18.638 | -102.3303
18.6417 | -102.3381
18.6454 | -102.344
18.6491 | -102.3503
18.6528 | -102.3583
18.6565 | -102.3676
18.6602 | -102.3758
18.6638 | -102.3842
18.6675 | -102.3937
18.6712 | -102.4031
18.6749 | -102.4108
18.6786 | -102.4172
18.6823 | -102.4254
18.686 | -102.4333
18.6897 | -102.4418
18.6933 | -102.4495
18.697 | -102.4568
18.7007 | -102.4653
18.7044 | -102.4726
18.7081 | -102.4822
18.7118 | -102.4881
18.7155 | -102.4926
18.7192 | -102.498
18.7229 | -102.5035
18.7265 | -102.5091
18.7302 | -102.512
18.7339 | -102.5144
18.7376 | -102.5175
18.7413 | -102.5207
18.745 | -102.5254
18.7487 | -102.5283
18.7524 | -102.5289
18.7561 | -102.5276
18.7597 | -102.5246
18.7634 | -102.518
18.7671 | -102.5111
18.7708 | -102.5057
18.7745 | -102.4989
18.7782 | -102.4915
18.7819 | -102.4858
18.7856 | -102.4799
18.7892 | -102.4702
18.7929 | -102.4602
18.7966 | -102.4499
18.8003 | -102.4408
18.804 | -102.43
18.8077 | -102.4183
18.8114 | -102.4052
18.8151 | -102.3935
18.8188 | -102.3811
18.8224 | -102.3686
18.8261 | -102.3553
18.8298 | -102.3416
18.8335 | -102.3258
18.8372 | -102.3108
18.8409 | -102.2974
18.8446 | -102.284
18.8483 | -102.272
18.852 | -102.2616
18.8556 | -102.2483
18.8593 | -102.2366
18.863 | -102.2281
18.8667 | -102.22
18.8704 | -102.2109
18.8741 | -102.202
18.8778 | -102.1934
18.8815 | -102.1876
18.8851 | -102.1808
18.8888 | -102.1747
18.8925 | -102.1716
18.8962 | -102.17
18.8999 | -102.1668
18.9036 | -102.1639
18.9073 | -102.1618
18.911 | -102.1601
18.9147 | -102.1592
18.9183 | -102.1595
18.922 | -102.1616
18.9257 | -102.1635
18.9294 | -102.168
18.9331 | -102.1714
18.9368 | -102.1736
18.9405 | -102.177
18.9442 | -102.1824
18.9479 | -102.1886
18.9515 | -102.1919
18.9552 | -102.1964
18.9589 | -102.2021
18.9626 | -102.2088
18.9663 | -102.2162
18.97 | -102.2225
18.9737 | -102.2316
18.9774 | -102.2415
18.981 | -102.2501
18.9847 | -102.257
18.9884 | -102.2652
18.9921 | -102.2727
18.9958 | -102.2817
18.9995 | -102.2932
19.0032 | -102.3034
19.0069 | -102.3117
19.0106 | -102.3207
19.0142 | -102.3299
19.0179 | -102.3388
19.0216 | -102.3504
19.0253 | -102.3631
19.029 | -102.3746
19.0327 | -102.3881
19.0364 | -102.3988
19.0401 | -102.4089
19.0438 | -102.4204
19.0474 | -102.4311
19.0511 | -102.4426
19.0548 | -102.452
19.0585 | -102.4611
19.0622 | -102.4696
19.0659 | -102.4802
19.0696 | -102.4919
19.0733 | -102.5013
19.077 | -102.5081
19.0806 | -102.5162
19.0843 | -102.5248
19.088 | -102.532
19.0917 | -102.5389
19.0954 | -102.5468
19.0991 | -102.5525
19.1028 | -102.556
19.1065 | -102.5548
19.1101 | -102.5538
19.1138 | -102.5536
19.1175 | -102.5515
19.1212 | -102.5464
19.1249 | -102.5396
19.1286 | -102.5321
19.1323 | -102.5202
19.136 | -102.5091
19.1397 | -102.4997
19.1433 | -102.4879
19.147 | -102.475
19.1507 | -102.4603
19.1544 | -102.4385
19.1581 | -102.4236
19.1618 | -102.4042
19.1655 | -102.3821
19.1692 | -102.3618
19.1729 | -102.3432
19.1765 | -102.3258
19.1802 | -102.3066
19.1839 | -102.2875
19.1876 | -102.268
19.1913 | -102.2467
19.195 | -102.2254
19.1987 | -102.207
19.2024 | -102.1886
19.206 | -102.1717
19.2097 | -102.1576
19.2134 | -102.1434
19.2171 | -102.1285
19.2208 | -102.1169
19.2245 | -102.1097
19.2282 | -102.1017
19.2319 | -102.0936
19.2356 | -102.0873
19.2392 | -102.0831
19.2429 | -102.0809
19.2466 | -102.0788
19.2503 | -102.081
19.254 | -102.0876
19.2577 | -102.0961
19.2614 | -102.1052
19.2651 | -102.1165
19.2688 | -102.1282
19.2724 | -102.1408
19.2761 | -102.157
19.2798 | -102.1766
19.2835 | -102.1953
19.2872 | -102.2166
19.2909 | -102.2391
19.2946 | -102.2596
19.2983 | -102.2803
19.302 | -102.3015
19.3056 | -102.3213
19.3093 | -102.3432
19.313 | -102.3649
19.3167 | -102.3854
19.3204 | -102.4056
19.3241 | -102.4241
19.3278 | -102.4427
19.3315 | -102.4621
19.3351 | -102.4806
19.3388 | -102.4957
19.3425 | -102.5076
19.3462 | -102.5226
19.3499 | -102.5376
19.3536 | -102.5492
19.3573 | -102.5589
19.361 | -102.5691
19.3647 | -102.5764
19.3683 | -102.5836
19.372 | -102.5874
19.3757 | -102.5929
19.3794 | -102.6005
19.3831 | -102.6038
19.3868 | -102.6049
19.3905 | -102.6057
19.3942 | -102.608
19.3979 | -102.6108
19.4015 | -102.6133
19.4052 | -102.6156
19.4089 | -102.6147
19.4126 | -102.6124
19.4163 | -102.6107
19.42 | -102.6074
19.4237 | -102.6051
19.4274 | -102.5997
19.431 | -102.5942
19.4347 | -102.5889
19.4384 | -102.5855
19.4421 | -102.5822
19.4458 | -102.5755
19.4495 | -102.568
19.4532 | -102.559
19.4569 | -102.5491
19.4606 | -102.5398
19.4642 | -102.5298
19.4679 | -102.5218
19.4716 | -102.5144
19.4753 | -102.5058
19.479 | -102.4957
19.4827 | -102.4858
19.4864 | -102.4766
19.4901 | -102.4665
19.4938 | -102.4577
19.4974 | -102.4489
19.5011 | -102.4407
19.5048 | -102.4342
19.5085 | -102.4284
19.5122 | -102.4245
19.5159 | -102.4214
19.5196 | -102.416
19.5233 | -102.4112
19.5269 | -102.4063
19.5306 | -102.4035
19.5343 | -102.3999
19.538 | -102.3981
19.5417 | -102.3981
19.5454 | -102.3978
19.5491 | -102.3987
19.5528 | -102.3996
19.5565 | -102.4015
19.5601 | -102.4072
19.5638 | -102.4119
19.5675 | -102.4158
19.5712 | -102.4231
19.5749 | -102.4309
19.5786 | -102.4401
19.5823 | -102.4495
19.586 | -102.4583
19.5897 | -102.4684
19.5933 | -102.4781
19.597 | -102.4893
19.6007 | -102.4999
19.6044 | -102.511
19.6081 | -102.5235
19.6118 | -102.5376
19.6155 | -102.5514
19.6192 | -102.564
19.6228 | -102.5756
19.6265 | -102.586
19.6302 | -102.5985
19.6339 | -102.6106
19.6376 | -102.6212
19.6413 | -102.6341
19.645 | -102.6448
19.6487 | -102.6566
19.6524 | -102.6682
19.656 | -102.6788
19.6597 | -102.6895
19.6634 | -102.7
19.6671 | -102.7114
19.6708 | -102.7214
19.6745 | -102.7288
19.6782 | -102.7334
19.6819 | -102.7394
19.6856 | -102.7444
19.6892 | -102.7504
19.6929 | -102.7566
19.6966 | -102.7616
19.7003 | -102.7658
19.704 | -102.7678
19.7077 | -102.7705
19.7114 | -102.7748
19.7151 | -102.7778
19.7188 | -102.7813
19.7224 | -102.7845
19.7261 | -102.7868
19.7298 | -102.7886
19.7335 | -102.7911
19.7372 | -102.795
19.7409 | -102.7977
19.7446 | -102.8022
19.7483 | -102.8066
19.7519 | -102.8087
19.7556 | -102.8103
19.7593 | -102.8134
19.763 | -102.818
19.7667 | -102.8207
19.7704 | -102.8248
19.7741 | -102.8285
19.7778 | -102.8311
19.7815 | -102.8354
19.7851 | -102.841
19.7888 | -102.8447
19.7925 | -102.8496
19.7962 | -102.8531
19.7999 | -102.8592
19.8036 | -102.8659
19.8073 | -102.8727
19.811 | -102.8781
19.8147 | -102.8853
19.8183 | -102.8907
19.822 | -102.8937
19.8257 | -102.9003
19.8294 | -102.9042
19.8331 | -102.9094
19.8368 | -102.9143
19.8405 | -102.9179
19.8442 | -102.922
19.8478 | -102.9272
19.8515 | -102.9323
19.8552 | -102.9389
19.8589 | -102.9445
19.8626 | -102.9512
19.8663 | -102.9587
19.87 | -102.966
19.8737 | -102.9756
19.8774 | -102.9845
19.881 | -102.9958
19.8847 | -103.0065
19.8884 | -103.0169
19.8921 | -103.028
19.8958 | -103.041
19.8995 | -103.0555
19.9032 | -103.0721
19.9069 | -103.0903
19.9106 | -103.1055
19.9142 | -103.1215
19.9179 | -103.14
19.9216 | -103.1574
19.9253 | -103.1776
19.929 | -103.1992
19.9327 | -103.2198
19.9364 | -103.2402
19.9401 | -103.2623
19.9438 | -103.2857
19.9474 | -103.3081
19.9511 | -103.33
19.9548 | -103.3537
19.9585 | -103.3788
19.9622 | -103.4012
19.9659 | -103.4226
19.9696 | -103.4454
19.9733 | -103.4671
19.9769 | -103.4881
19.9806 | -103.5106
19.9843 | -103.5334
19.988 | -103.5537
19.9917 | -103.5735
19.9954 | -103.5913
19.9991 | -103.6088
20.0028 | -103.6244
20.0065 | -103.6401
20.0101 | -103.6556
20.0138 | -103.6704
20.0175 | -103.683
20.0212 | -103.6934
20.0249 | -103.7048
20.0286 | -103.7153
20.0323 | -103.7233
20.036 | -103.7327
20.0397 | -103.7416
20.0433 | -103.7472
20.047 | -103.7507
20.0507 | -103.7566
20.0544 | -103.7612
20.0581 | -103.7643
20.0618 | -103.7682
20.0655 | -103.7666
20.0692 | -103.7638
20.0728 | -103.7643
20.0765 | -103.7616
20.0802 | -103.7596
20.0839 | -103.7559
20.0876 | -103.752
20.0913 | -103.7467
20.095 | -103.7421
20.0987 | -103.7375
20.1024 | -103.7317
20.106 | -103.7275
20.1097 | -103.7218
20.1134 | -103.7196
20.1171 | -103.7154
20.1208 | -103.709
20.1245 | -103.7036
20.1282 | -103.6983
20.1319 | -103.6935
20.1356 | -103.6897
20.1392 | -103.6862
20.1429 | -103.6838
20.1466 | -103.6821
20.1503 | -103.6798
20.154 | -103.6766
20.1577 | -103.6732
20.1614 | -103.6704
20.1651 | -103.6685
20.1687 | -103.6688
20.1724 | -103.6696
20.1761 | -103.6705
20.1798 | -103.6717
20.1835 | -103.6734
20.1872 | -103.6753
20.1909 | -103.679
20.1946 | -103.6851
20.1983 | -103.6906
20.2019 | -103.6951
20.2056 | -103.6994
20.2093 | -103.7033
20.213 | -103.7084
20.2167 | -103.7139
20.2204 | -103.7198
20.2241 | -103.7239
20.2278 | -103.7285
20.2315 | -103.7318
20.2351 | -103.7371
20.2388 | -103.7444
20.2425 | -103.754
20.2462 | -103.763
20.2499 | -103.7707
20.2536 | -103.7774
20.2573 | -103.7838
20.261 | -103.7908
20.2646 | -103.7985
20.2683 | -103.8047
20.272 | -103.8118
20.2757 | -103.8208
20.2794 | -103.8293
20.2831 | -103.8366
20.2868 | -103.8429
20.2905 | -103.849
20.2942 | -103.8569
20.2978 | -103.8647
20.3015 | -103.8713
20.3052 | -103.8792
20.3089 | -103.8884
20.3126 | -103.8982
20.3163 | -103.9073
20.32 | -103.9161
20.3237 | -103.9254
20.3274 | -103.9357
20.331 | -103.9462
20.3347 | -103.9541
20.3384 | -103.9624
20.3421 | -103.9723
20.3458 | -103.9819
20.3495 | -103.9878
20.3532 | -103.9968
20.3569 | -104.0059
20.3606 | -104.0118
20.3642 | -104.0184
20.3679 | -104.0231
20.3716 | -104.029
20.3753 | -104.0346
20.379 | -104.0379
20.3827 | -104.0406
20.3864 | -104.0416
20.3901 | -104.0451
20.3937 | -104.0455
20.3974 | -104.0455
20.4011 | -104.044
20.4048 | -104.0428
20.4085 | -104.0417
20.4122 | -104.0371
20.4159 | -104.0305
20.4196 | -104.026
20.4233 | -104.0209
20.4269 | -104.0144
20.4306 | -104.0055
20.4343 | -103.9984
20.438 | -103.9925
20.4417 | -103.985
20.4454 | -103.9787
20.4491 | -103.9721
20.4528 | -103.9638
20.4565 | -103.956
20.4601 | -103.9476
20.4638 | -103.9394
20.4675 | -103.9318
20.4712 | -103.9245
20.4749 | -103.9146
20.4786 | -103.9061
20.4823 | -103.8953
20.486 | -103.8851
20.4896 | -103.877
20.4933 | -103.8702
20.497 | -103.8641
20.5007 | -103.8584
20.5044 | -103.8513
20.5081 | -103.8468
20.5118 | -103.841
20.5155 | -103.837
20.5192 | -103.8302
20.5228 | -103.8223
20.5265 | -103.8136
20.5302 | -103.8053
20.5339 | -103.7936
20.5376 | -103.7801
20.5413 | -103.7705
20.545 | -103.7595
20.5487 | -103.7446
20.5524 | -103.7325
20.556 | -103.7217
20.5597 | -103.7102
20.5634 | -103.6941
20.5671 | -103.6779
20.5708 | -103.6623
20.5745 | -103.6456
20.5782 | -103.6306
20.5819 | -103.6132
20.5856 | -103.5899
20.5892 | -103.5679
20.5929 | -103.5432
20.5966 | -103.5122
20.6003 | -103.4824
20.604 | -103.448
20.6077 | -103.4076
20.6114 | -103.3695
20.6151 | -103.3323
20.6187 | -103.2954
20.6224 | -103.2559
20.6261 | -103.213
20.6298 | -103.1668
20.6335 | -103.1191
20.6372 | -103.0728
20.6409 | -103.0283
20.6446 | -102.9801
20.6483 | -102.9256
20.6519 | -102.8717
20.6556 | -102.8189
20.6593 | -102.7635
20.663 | -102.7076
20.6667 | -102.653
20.6704 | -102.5997
20.6741 | -102.5461
20.6778 | -102.4927
20.6815 | -102.4386
20.6851 | -102.3872
20.6888 | -102.3368
20.6925 | -102.2864
20.6962 | -102.2379
20.6999 | -102.1927
20.7036 | -102.1501
20.7073 | -102.1087
20.711 | -102.0667
20.7146 | -102.0231
20.7183 | -101.9821
20.722 | -101.9454
20.7257 | -101.9119
20.7294 | -101.8819
20.7331 | -101.8544
20.7368 | -101.8284
20.7405 | -101.8056
20.7442 | -101.7872
20.7478 | -101.7691
20.7515 | -101.7676
20.7552 | -101.753
20.7589 | -101.7402
20.7626 | -101.731
20.7663 | -101.7229
20.77 | -101.7175
20.7737 | -101.7109
20.7774 | -101.7089
20.781 | -101.7101
20.7847 | -101.7111
20.7884 | -101.7117
20.7921 | -101.714
20.7958 | -101.7189
20.7995 | -101.7259
20.8032 | -101.7331
20.8069 | -101.7402
20.8105 | -101.7469
20.8142 | -101.7561
20.8179 | -101.7687
20.8216 | -101.7806
20.8253 | -101.7911
20.829 | -101.8038
20.8327 | -101.8168
20.8364 | -101.8291
20.8401 | -101.844
20.8437 | -101.8577
20.8474 | -101.8675
20.8511 | -101.8766
20.8548 | -101.8866
20.8585 | -101.8974
20.8622 | -101.9061
20.8659 | -101.9122
20.8696 | -101.9197
20.8733 | -101.9278
20.8769 | -101.9348
20.8806 | -101.9415
20.8843 | -101.9458
20.888 | -101.9497
20.8917 | -101.954
20.8954 | -101.9552
20.8991 | -101.9564
20.9028 | -101.9601
20.9064 | -101.9619
20.9101 | -101.9612
20.9138 | -101.9581
20.9175 | -101.957
20.9212 | -101.9564
20.9249 | -101.9558
20.9286 | -101.9539
20.9323 | -101.9523
20.936 | -101.9496
20.9396 | -101.9486
20.9433 | -101.9525
20.947 | -101.9562
20.9507 | -101.9584
20.9544 | -101.9625
20.9581 | -101.9647
20.9618 | -101.9654
20.9655 | -101.9678
20.9692 | -101.9715
20.9728 | -101.9762
20.9765 | -101.9828
20.9802 | -101.9874
20.9839 | -101.9932
20.9876 | -102.0015
20.9913 | -102.0073
20.995 | -102.0116
20.9987 | -102.0179
21.0024 | -102.0264
21.006 | -102.0364
21.0097 | -102.051
21.0134 | -102.0637
21.0171 | -102.074
21.0208 | -102.0827
21.0245 | -102.0915
21.0282 | -102.1017
21.0319 | -102.1124
21.0355 | -102.123
21.0392 | -102.1318
21.0429 | -102.1409
21.0466 | -102.1516
21.0503 | -102.1641
21.054 | -102.1766
21.0577 | -102.1885
21.0614 | -102.1997
21.0651 | -102.2115
21.0687 | -102.223
21.0724 | -102.2387
21.0761 | -102.2524
21.0798 | -102.2631
21.0835 | -102.2741
21.0872 | -102.2873
21.0909 | -102.3009
21.0946 | -102.3109
21.0983 | -102.319
21.1019 | -102.3281
21.1056 | -102.3395
21.1093 | -102.3528
21.113 | -102.3634
21.1167 | -102.3719
21.1204 | -102.3796
21.1241 | -102.3888
21.1278 | -102.3972
21.1314 | -102.4066
21.1351 | -102.4146
21.1388 | -102.4198
21.1425 | -102.4268
21.1462 | -102.4337
21.1499 | -102.4382
21.1536 | -102.4429
21.1573 | -102.4479
21.161 | -102.4521
21.1646 | -102.4565
21.1683 | -102.4588
21.172 | -102.458
21.1757 | -102.459
21.1794 | -102.4608
21.1831 | -102.4631
21.1868 | -102.4676
21.1905 | -102.4712
21.1942 | -102.4727
21.1978 | -102.4737
21.2015 | -102.4742
21.2052 | -102.4717
21.2089 | -102.4696
21.2126 | -102.4676
21.2163 | -102.4643
21.22 | -102.4643
21.2237 | -102.4628
21.2274 | -102.4604
21.231 | -102.4567
21.2347 | -102.4546
21.2384 | -102.4524
21.2421 | -102.4492
21.2458 | -102.4459
21.2495 | -102.4421
21.2532 | -102.4384
21.2569 | -102.4379
21.2605 | -102.4376
21.2642 | -102.4353
21.2679 | -102.4333
21.2716 | -102.432
21.2753 | -102.4312
21.279 | -102.4328
21.2827 | -102.4392
21.2864 | -102.4421
21.2901 | -102.4435
21.2937 | -102.4481
21.2974 | -102.4537
21.3011 | -102.4586
21.3048 | -102.464
21.3085 | -102.4698
21.3122 | -102.4747
21.3159 | -102.4815
21.3196 | -102.4897
21.3233 | -102.496
21.3269 | -102.5004
21.3306 | -102.5077
21.3343 | -102.5148
21.338 | -102.5236
21.3417 | -102.5317
21.3454 | -102.5404
21.3491 | -102.55
21.3528 | -102.555
21.3564 | -102.5562
21.3601 | -102.5598
21.3638 | -102.5643
21.3675 | -102.5688
21.3712 | -102.5725
21.3749 | -102.5739
21.3786 | -102.574
21.3823 | -102.5733
21.386 | -102.5742
21.3896 | -102.574
21.3933 | -102.5704
21.397 | -102.5669
21.4007 | -102.5639
21.4044 | -102.5594
21.4081 | -102.5524
21.4118 | -102.5424
21.4155 | -102.5302
21.4192 | -102.5142
21.4228 | -102.4994
21.4265 | -102.4849
21.4302 | -102.4685
21.4339 | -102.4519
21.4376 | -102.4358
21.4413 | -102.4183
21.445 | -102.4006
21.4487 | -102.3828
21.4523 | -102.3649
21.456 | -102.3453
21.4597 | -102.3253
21.4634 | -102.3058
21.4671 | -102.2855
21.4708 | -102.2659
21.4745 | -102.2482
21.4782 | -102.2287
21.4819 | -102.2079
21.4855 | -102.1859
21.4892 | -102.1641
21.4929 | -102.1453
21.4966 | -102.1277
21.5003 | -102.1075
21.504 | -102.0869
21.5077 | -102.0676
21.5114 | -102.0465
21.5151 | -102.0275
21.5187 | -102.0099
21.5224 | -101.9912
21.5261 | -101.9721
21.5298 | -101.9529
21.5335 | -101.9339
21.5372 | -101.914
21.5409 | -101.8959
21.5446 | -101.8757
21.5482 | -101.8557
21.5519 | -101.8381
21.5556 | -101.8209
21.5593 | -101.8043
21.563 | -101.7884
21.5667 | -101.7747
21.5704 | -101.7593
21.5741 | -101.7431
21.5778 | -101.7262
21.5814 | -101.7113
21.5851 | -101.6943
21.5888 | -101.6791
21.5925 | -101.6601
21.5962 | -101.6467
21.5999 | -101.6326
21.6036 | -101.6191
21.6073 | -101.6077
21.611 | -101.5971
21.6146 | -101.5858
21.6183 | -101.5764
21.622 | -101.5682
21.6257 | -101.5604
21.6294 | -101.5548
21.6331 | -101.5524
21.6368 | -101.5482
21.6405 | -101.5448
21.6442 | -101.5413
21.6478 | -101.5416
21.6515 | -101.5424
21.6552 | -101.5415
21.6589 | -101.5447
21.6626 | -101.5522
21.6663 | -101.557
21.67 | -101.5622
21.6737 | -101.569
21.6773 | -101.5758
21.681 | -101.5831
21.6847 | -101.5937
21.6884 | -101.6069
21.6921 | -101.6192
21.6958 | -101.6329
21.6995 | -101.6456
21.7032 | -101.6599
21.7069 | -101.6769
21.7105 | -101.6951
21.7142 | -101.7138
21.7179 | -101.7325
21.7216 | -101.7544
21.7253 | -101.7788
21.729 | -101.8028
21.7327 | -101.825
21.7364 | -101.8449
21.7401 | -101.8644
21.7437 | -101.8831
21.7474 | -101.9037
21.7511 | -101.9238
21.7548 | -101.9451
21.7585 | -101.9655
21.7622 | -101.9862
21.7659 | -102.0059
21.7696 | -102.0249
21.7732 | -102.0436
21.7769 | -102.0582
21.7806 | -102.0681
21.7843 | -102.0805
21.788 | -102.0968
21.7917 | -102.1115
21.7954 | -102.1227
21.7991 | -102.1296
21.8028 | -102.135
21.8064 | -102.1433
21.8101 | -102.1511
21.8138 | -102.1582
21.8175 | -102.1608
21.8212 | -102.168
21.8249 | -102.1733
21.8286 | -102.1769
21.8323 | -102.1809
21.836 | -102.1852
21.8396 | -102.1901
21.8433 | -102.1944
21.847 | -102.198
21.8507 | -102.1998
21.8544 | -102.2039
21.8581 | -102.2065
21.8618 | -102.2066
21.8655 | -102.2081
21.8692 | -102.2073
21.8728 | -102.2066
21.8765 | -102.2059
21.8802 | -102.2072
21.8839 | -102.2099
21.8876 | -102.2139
21.8913 | -102.2156
21.895 | -102.2208
21.8987 | -102.2243
21.9023 | -102.2302
21.906 | -102.2323
21.9097 | -102.2373
21.9134 | -102.242
21.9171 | -102.2444
21.9208 | -102.2489
21.9245 | -102.2564
21.9282 | -102.2613
21.9319 | -102.2666
21.9355 | -102.274
21.9392 | -102.2827
21.9429 | -102.2902
21.9466 | -102.2982
21.9503 | -102.3031
21.954 | -102.3111
21.9577 | -102.3212
21.9614 | -102.3321
21.9651 | -102.3438
21.9687 | -102.3554
21.9724 | -102.3658
21.9761 | -102.378
21.9798 | -102.3877
21.9835 | -102.3957
21.9872 | -102.4075
21.9909 | -102.4204
21.9946 | -102.4318
21.9982 | -102.4422
22.0019 | -102.4494
22.0056 | -102.4582
22.0093 | -102.4724
22.013 | -102.4849
22.0167 | -102.4946
22.0204 | -102.5044
22.0241 | -102.5115
22.0278 | -102.5197
22.0314 | -102.5292
22.0351 | -102.5348
22.0388 | -102.5385
22.0425 | -102.5442
22.0462 | -102.5518
22.0499 | -102.5595
22.0536 | -102.5685
22.0573 | -102.5745
22.061 | -102.5821
22.0646 | -102.5869
22.0683 | -102.5922
22.072 | -102.5975
22.0757 | -102.6052
22.0794 | -102.6124
22.0831 | -102.6174
22.0868 | -102.624
22.0905 | -102.636
22.0941 | -102.6467
22.0978 | -102.6536
22.1015 | -102.6618
22.1052 | -102.6717
22.1089 | -102.6845
22.1126 | -102.6976
22.1163 | -102.7092
22.12 | -102.7198
22.1237 | -102.7298
22.1273 | -102.7396
22.131 | -102.7501
22.1347 | -102.7623
22.1384 | -102.7754
22.1421 | -102.7894
22.1458 | -102.8018
22.1495 | -102.8162
22.1532 | -102.8318
22.1569 | -102.8478
22.1605 | -102.8661
22.1642 | -102.8843
22.1679 | -102.9032
22.1716 | -102.9237
22.1753 | -102.9445
22.179 | -102.9643
22.1827 | -102.9827
22.1864 | -103
22.19 | -103.0166
22.1937 | -103.0352
22.1974 | -103.0544
22.2011 | -103.0722
22.2048 | -103.0931
22.2085 | -103.1142
22.2122 | -103.1367
22.2159 | -103.1606
22.2196 | -103.1838
22.2232 | -103.2061
22.2269 | -103.2277
22.2306 | -103.2502
22.2343 | -103.2708
22.238 | -103.2904
22.2417 | -103.3131
22.2454 | -103.3372
22.2491 | -103.361
22.2528 | -103.3823
22.2564 | -103.4045
22.2601 | -103.4285
22.2638 | -103.4568
22.2675 | -103.4796
22.2712 | -103.4996
22.2749 | -103.5211
22.2786 | -103.544
22.2823 | -103.5662
22.286 | -103.5889
22.2896 | -103.6126
22.2933 | -103.6367
22.297 | -103.6636
22.3007 | -103.6881
22.3044 | -103.711
22.3081 | -103.732
22.3118 | -103.7545
22.3155 | -103.7838
22.3191 | -103.8093
22.3228 | -103.8322
22.3265 | -103.8574
22.3302 | -103.8835
22.3339 | -103.9083
22.3376 | -103.939
22.3413 | -103.9698
22.345 | -103.9969
22.3487 | -104.0249
22.3523 | -104.0559
22.356 | -104.0874
22.3597 | -104.1164
22.3634 | -104.1481
22.3671 | -104.178
22.3708 | -104.2057
22.3745 | -104.2356
22.3782 | -104.264
22.3819 | -104.2933
22.3855 | -104.322
22.3892 | -104.35
22.3929 | -104.3772
22.3966 | -104.4045
22.4003 | -104.4349
22.404 | -104.466
22.4077 | -104.4909
22.4114 | -104.5126
22.415 | -104.5337
22.4187 | -104.5549
22.4224 | -104.5749
22.4261 | -104.5932
22.4298 | -104.608
22.4335 | -104.6245
22.4372 | -104.6405
22.4409 | -104.6568
22.4446 | -104.6694
22.4482 | -104.6784
22.4519 | -104.6847
22.4556 | -104.686
22.4593 | -104.6847
22.463 | -104.6817
22.4667 | -104.6816
22.4704 | -104.6787
22.4741 | -104.6697
22.4778 | -104.6587
22.4814 | -104.648
22.4851 | -104.6374
22.4888 | -104.6258
22.4925 | -104.6116
22.4962 | -104.596
22.4999 | -104.5838
22.5036 | -104.571
22.5073 | -104.5531
22.511 | -104.5296
22.5146 | -104.5033
22.5183 | -104.4785
22.522 | -104.4523
22.5257 | -104.425
22.5294 | -104.3964
22.5331 | -104.3653
22.5368 | -104.3318
22.5405 | -104.2979
22.5441 | -104.2671
22.5478 | -104.2375
22.5515 | -104.2066
22.5552 | -104.1747
22.5589 | -104.1436
22.5626 | -104.1126
22.5663 | -104.0803
22.57 | -104.0492
22.5737 | -104.0194
22.5773 | -103.9915
22.581 | -103.9653
22.5847 | -103.9403
22.5884 | -103.9152
22.5921 | -103.8933
22.5958 | -103.8738
22.5995 | -103.8551
22.6032 | -103.8384
22.6069 | -103.8245
22.6105 | -103.8142
22.6142 | -103.8047
22.6179 | -103.7949
22.6216 | -103.789
22.6253 | -103.7872
22.629 | -103.787
22.6327 | -103.7866
22.6364 | -103.7881
22.64 | -103.7922
22.6437 | -103.7977
22.6474 | -103.8035
22.6511 | -103.8124
22.6548 | -103.8215
22.6585 | -103.8305
22.6622 | -103.8428
22.6659 | -103.8589
22.6696 | -103.8753
22.6732 | -103.8914
22.6769 | -103.9085
22.6806 | -103.9256
22.6843 | -103.9412
22.688 | -103.9577
22.6917 | -103.9739
22.6954 | -103.9909
22.6991 | -104.0081
22.7028 | -104.0252
22.7064 | -104.042
22.7101 | -104.0585
22.7138 | -104.0724
22.7175 | -104.0864
22.7212 | -104.0995
22.7249 | -104.1116
22.7286 | -104.1247
22.7323 | -104.1372
22.7359 | -104.1472
22.7396 | -104.1567
22.7433 | -104.1646
22.747 | -104.1703
22.7507 | -104.1789
22.7544 | -104.1852
22.7581 | -104.19
22.7618 | -104.1911
22.7655 | -104.193
22.7691 | -104.1955
22.7728 | -104.1982
22.7765 | -104.2018
22.7802 | -104.2052
22.7839 | -104.205
22.7876 | -104.1992
22.7913 | -104.1944
22.795 | -104.1929
22.7987 | -104.1888
22.8023 | -104.1844
22.806 | -104.1766
22.8097 | -104.1679
22.8134 | -104.1599
22.8171 | -104.1516
22.8208 | -104.1427
22.8245 | -104.1352
22.8282 | -104.1303
22.8318 | -104.125
22.8355 | -104.1214
22.8392 | -104.1148
22.8429 | -104.106
22.8466 | -104.0971
22.8503 | -104.088
22.854 | -104.0824
22.8577 | -104.0797
22.8614 | -104.0771
22.865 | -104.0722
22.8687 | -104.0632
22.8724 | -104.0538
22.8761 | -104.0455
22.8798 | -104.0386
22.8835 | -104.0338
22.8872 | -104.029
22.8909 | -104.0233
22.8946 | -104.0164
22.8982 | -104.01
22.9019 | -104.0026
22.9056 | -103.9943
22.9093 | -103.9874
22.913 | -103.9804
22.9167 | -103.9728
22.9204 | -103.9639
22.9241 | -103.9542
22.9278 | -103.9456
22.9314 | -103.9375
22.9351 | -103.9293
22.9388 | -103.9207
22.9425 | -103.9136
22.9462 | -103.9061
22.9499 | -103.8963
22.9536 | -103.884
22.9573 | -103.8731
22.9609 | -103.8649
22.9646 | -103.8546
22.9683 | -103.8434
22.972 | -103.8364
22.9757 | -103.8269
22.9794 | -103.8166
22.9831 | -103.809
22.9868 | -103.8011
22.9905 | -103.7918
22.9941 | -103.7822
22.9978 | -103.7732
23.0015 | -103.7638
23.0052 | -103.7536
23.0089 | -103.7443
23.0126 | -103.735
23.0163 | -103.727
23.02 | -103.7185
23.0237 | -103.7084
23.0273 | -103.6991
23.031 | -103.6915
23.0347 | -103.6836
23.0384 | -103.6767
23.0421 | -103.6704
23.0458 | -103.6656
23.0495 | -103.6612
23.0532 | -103.6551
23.0568 | -103.6499
23.0605 | -103.6471
23.0642 | -103.6432
23.0679 | -103.6425
23.0716 | -103.6429
23.0753 | -103.641
23.079 | -103.6374
23.0827 | -103.6336
23.0864 | -103.6343
23.09 | -103.6346
23.0937 | -103.6351
23.0974 | -103.6357
23.1011 | -103.6368
23.1048 | -103.6364
23.1085 | -103.6383
23.1122 | -103.6413
23.1159 | -103.6435
23.1196 | -103.6443
23.1232 | -103.6457
23.1269 | -103.6459
23.1306 | -103.6495
23.1343 | -103.6539
23.138 | -103.6577
23.1417 | -103.6627
23.1454 | -103.6675
23.1491 | -103.6707
23.1528 | -103.6737
23.1564 | -103.6782
23.1601 | -103.6823
23.1638 | -103.6881
23.1675 | -103.6948
23.1712 | -103.7006
23.1749 | -103.7081
23.1786 | -103.7148
23.1823 | -103.7195
23.1859 | -103.724
23.1896 | -103.7303
23.1933 | -103.7394
23.197 | -103.7478
23.2007 | -103.7579
23.2044 | -103.7653
23.2081 | -103.7714
23.2118 | -103.777
23.2155 | -103.7832
23.2191 | -103.7919
23.2228 | -103.8026
23.2265 | -103.811
23.2302 | -103.8202
23.2339 | -103.832
23.2376 | -103.8431
23.2413 | -103.8535
23.245 | -103.8626
23.2487 | -103.8749
23.2523 | -103.8882
23.256 | -103.899
23.2597 | -103.909
23.2634 | -103.9185
23.2671 | -103.93
23.2708 | -103.9414
23.2745 | -103.9517
23.2782 | -103.9647
23.2818 | -103.9804
23.2855 | -103.9945
23.2892 | -104.0065
23.2929 | -104.0189
23.2966 | -104.0316
23.3003 | -104.0415
23.304 | -104.0524
23.3077 | -104.0657
23.3114 | -104.0779
23.315 | -104.0896
23.3187 | -104.1012
23.3224 | -104.1135
23.3261 | -104.1281
23.3298 | -104.1438
23.3335 | -104.1593
23.3372 | -104.1737
23.3409 | -104.1859
23.3446 | -104.1961
23.3482 | -104.2088
23.3519 | -104.2211
23.3556 | -104.2335
23.3593 | -104.2448
23.363 | -104.2568
23.3667 | -104.2691
23.3704 | -104.2812
23.3741 | -104.2915
23.3777 | -104.3007
23.3814 | -104.3113
23.3851 | -104.322
23.3888 | -104.3305
23.3925 | -104.3387
23.3962 | -104.347
23.3999 | -104.3565
23.4036 | -104.3659
23.4073 | -104.3764
23.4109 | -104.3803
23.4146 | -104.3874
23.4183 | -104.3926
23.422 | -104.3969
23.4257 | -104.4003
23.4294 | -104.4028
23.4331 | -104.404
23.4368 | -104.4072
23.4405 | -104.4082
23.4441 | -104.4062
23.4478 | -104.4053
23.4515 | -104.4046
23.4552 | -104.4031
23.4589 | -104.4019
23.4626 | -104.3968
23.4663 | -104.3912
23.47 | -104.3854
23.4736 | -104.3779
23.4773 | -104.3697
23.481 | -104.3636
23.4847 | -104.3544
23.4884 | -104.3445
23.4921 | -104.336
23.4958 | -104.3267
23.4995 | -104.317
23.5032 | -104.3072
23.5068 | -104.2933
23.5105 | -104.2803
23.5142 | -104.2701
23.5179 | -104.26
23.5216 | -104.2501
23.5253 | -104.2385
23.529 | -104.227
23.5327 | -104.2148
23.5364 | -104.204
23.54 | -104.1931
23.5437 | -104.18
23.5474 | -104.1676
23.5511 | -104.155
23.5548 | -104.1428
23.5585 | -104.1322
23.5622 | -104.1215
23.5659 | -104.1095
23.5696 | -104.0977
23.5732 | -104.0875
23.5769 | -104.0782
23.5806 | -104.0684
23.5843 | -104.059
23.588 | -104.0502
23.5917 | -104.0421
23.5954 | -104.0344
23.5991 | -104.0256
23.6027 | -104.0183
23.6064 | -104.0109
23.6101 | -104.0028
23.6138 | -103.9953
23.6175 | -103.9872
23.6212 | -103.9786
23.6249 | -103.9712
23.6286 | -103.9643
23.6323 | -103.9574
23.6359 | -103.9511
23.6396 | -103.945
23.6433 | -103.9394
23.647 | -103.9342
23.6507 | -103.9271
23.6544 | -103.9209
23.6581 | -103.9151
23.6618 | -103.9093
23.6655 | -103.904
23.6691 | -103.8985
23.6728 | -103.8951
23.6765 | -103.8914
23.6802 | -103.8861
23.6839 | -103.8806
23.6876 | -103.8742
23.6913 | -103.8676
23.695 | -103.8632
23.6986 | -103.8592
23.7023 | -103.854
23.706 | -103.8492
23.7097 | -103.8437
23.7134 | -103.8376
23.7171 | -103.8326
23.7208 | -103.8284
23.7245 | -103.8235
23.7282 | -103.818
23.7318 | -103.812
23.7355 | -103.806
23.7392 | -103.8014
23.7429 | -103.7965
23.7466 | -103.7932
23.7503 | -103.7915
23.754 | -103.7875
23.7577 | -103.7815
23.7614 | -103.7771
23.765 | -103.7745
23.7687 | -103.7703
23.7724 | -103.7647
23.7761 | -103.7591
23.7798 | -103.7543
23.7835 | -103.7493
23.7872 | -103.7445
23.7909 | -103.7407
23.7946 | -103.7338
23.7982 | -103.7282
23.8019 | -103.7236
23.8056 | -103.7194
23.8093 | -103.7148
23.813 | -103.7084
23.8167 | -103.7007
23.8204 | -103.6967
23.8241 | -103.6928
23.8277 | -103.6856
23.8314 | -103.6794
23.8351 | -103.6718
23.8388 | -103.6648
23.8425 | -103.6614
23.8462 | -103.6555
23.8499 | -103.6477
23.8536 | -103.6418
23.8573 | -103.6357
23.8609 | -103.6294
23.8646 | -103.6254
23.8683 | -103.6202
23.872 | -103.6142
23.8757 | -103.6098
23.8794 | -103.6052
23.8831 | -103.5998
23.8868 | -103.5942
23.8905 | -103.5882
23.8941 | -103.5828
23.8978 | -103.5774
23.9015 | -103.5712
23.9052 | -103.5652
23.9089 | -103.5597
23.9126 | -103.5543
23.9163 | -103.5493
23.92 | -103.5444
23.9236 | -103.5384
23.9273 | -103.533
23.931 | -103.5289
23.9347 | -103.5252
23.9384 | -103.5195
23.9421 | -103.5158
23.9458 | -103.5138
23.9495 | -103.5108
23.9532 | -103.5075
23.9568 | -103.505
23.9605 | -103.5052
23.9642 | -103.5051
23.9679 | -103.5047
23.9716 | -103.5048
23.9753 | -103.5068
23.979 | -103.5077
23.9827 | -103.5069
23.9864 | -103.5053
23.99 | -103.5047
23.9937 | -103.5047
23.9974 | -103.5044
24.0011 | -103.5057
24.0048 | -103.5075
24.0085 | -103.5083
24.0122 | -103.5095
24.0159 | -103.5118
24.0195 | -103.5133
24.0232 | -103.5174
24.0269 | -103.5208
24.0306 | -103.5241
24.0343 | -103.526
24.038 | -103.5278
24.0417 | -103.5302
24.0454 | -103.531
24.0491 | -103.5316
24.0527 | -103.5329
24.0564 | -103.5303
24.0601 | -103.527
24.0638 | -103.5283
24.0675 | -103.529
24.0712 | -103.5284
24.0749 | -103.5288
24.0786 | -103.5279
24.0823 | -103.5262
24.0859 | -103.5253
24.0896 | -103.5228
24.0933 | -103.5218
24.097 | -103.5222
24.1007 | -103.5192
24.1044 | -103.5145
24.1081 | -103.5115
24.1118 | -103.5084
24.1154 | -103.5054
24.1191 | -103.5011
24.1228 | -103.4965
24.1265 | -103.4911
24.1302 | -103.4845
24.1339 | -103.4804
24.1376 | -103.4756
24.1413 | -103.4701
24.145 | -103.4641
24.1486 | -103.4594
24.1523 | -103.457
24.156 | -103.4539
24.1597 | -103.4492
24.1634 | -103.4431
24.1671 | -103.4393
24.1708 | -103.4343
24.1745 | -103.4286
24.1782 | -103.4241
24.1818 | -103.4191
24.1855 | -103.4147
24.1892 | -103.4094
24.1929 | -103.3994
24.1966 | -103.3892
24.2003 | -103.3835
24.204 | -103.3765
24.2077 | -103.3697
24.2114 | -103.3639
24.215 | -103.3567
24.2187 | -103.3498
24.2224 | -103.3439
24.2261 | -103.3386
24.2298 | -103.3333
24.2335 | -103.326
24.2372 | -103.3193
24.2409 | -103.3122
24.2445 | -103.3059
24.2482 | -103.2997
24.2519 | -103.2925
24.2556 | -103.2839
24.2593 | -103.2767
24.263 | -103.2692
24.2667 | -103.2592
24.2704 | -103.2415
24.2741 | -103.2055
24.2777 | -103.1541
24.2814 | -103.0899
24.2851 | -103.0335
24.2888 | -102.9973
24.2925 | -102.9623
24.2962 | -102.9258
24.2999 | -102.888
24.3036 | -102.8474
24.3073 | -102.8024
24.3109 | -102.7568
24.3146 | -102.7102
24.3183 | -102.6602
24.322 | -102.609
24.3257 | -102.5577
24.3294 | -102.5047
24.3331 | -102.4509
24.3368 | -102.3952
24.3404 | -102.3358
24.3441 | -102.2754
24.3478 | -102.2146
24.3515 | -102.1541
24.3552 | -102.0927
24.3589 | -102.0305
24.3626 | -101.967
24.3663 | -101.9028
24.37 | -101.8404
24.3736 | -101.7783
24.3773 | -101.7167
24.381 | -101.6566
24.3847 | -101.5979
24.3884 | -101.5402
24.3921 | -101.4863
24.3958 | -101.4344
24.3995 | -101.3862
24.4032 | -101.3415
24.4068 | -101.2975
24.4105 | -101.2557
24.4142 | -101.2185
24.4179 | -101.1849
24.4216 | -101.1532
24.4253 | -101.1266
24.429 | -101.1045
24.4327 | -101.0872
24.4364 | -101.074
24.44 | -101.0648
24.4437 | -101.0601
24.4474 | -101.0595
24.4511 | -101.0634
24.4548 | -101.0704
24.4585 | -101.0812
24.4622 | -101.0978
24.4659 | -101.1191
24.4695 | -101.1442
24.4732 | -101.1741
24.4769 | -101.234
24.4806 | -101.2691
24.4843 | -101.3071
24.488 | -101.3483
24.4917 | -101.3934
24.4954 | -101.4414
24.4991 | -101.4912
24.5027 | -101.5429
24.5064 | -101.5966
24.5101 | -101.6531
24.5138 | -101.7119
24.5175 | -101.7711
24.5212 | -101.8316
24.5249 | -101.8926
24.5286 | -101.9531
24.5323 | -102.0146
24.5359 | -102.0769
24.5396 | -102.1414
24.5433 | -102.2044
24.547 | -102.2612
24.5507 | -102.3145
24.5544 | -102.3636
24.5581 | -102.4117
24.5618 | -102.4569
24.5654 | -102.4961
24.5691 | -102.5318
24.5728 | -102.5593
24.5765 | -102.5761
24.5802 | -102.5733
24.5839 | -102.5481
24.5876 | -102.5082
24.5913 | -102.486
24.595 | -102.4697
24.5986 | -102.4479
24.6023 | -102.4224
24.606 | -102.3936
24.6097 | -102.3589
24.6134 | -102.3188
24.6171 | -102.2784
24.6208 | -102.2439
24.6245 | -102.2123
24.6282 | -102.1772
24.6318 | -102.1234
24.6355 | -102.0462
24.6392 | -101.9637
24.6429 | -101.8781
24.6466 | -101.7914
24.6503 | -101.7034
24.654 | -101.616
24.6577 | -101.5261
24.6613 | -101.4351
24.665 | -101.3432
24.6687 | -101.2518
24.6724 | -101.1612
24.6761 | -101.072
24.6798 | -100.9831
24.6835 | -100.8964
24.6872 | -100.8122
24.6909 | -100.7296
24.6945 | -100.6494
24.6982 | -100.5721
24.7019 | -100.4986
24.7056 | -100.4287
24.7093 | -100.3635
24.713 | -100.3027
24.7167 | -100.2452
24.7204 | -100.1921
24.7241 | -100.1443
24.7277 | -100.1021
24.7314 | -100.0654
24.7351 | -100.0351
24.7388 | -100.0091
24.7425 | -99.9874
24.7462 | -99.9714
24.7499 | -99.9607
24.7536 | -99.955
24.7572 | -99.9552
24.7609 | -99.9614
24.7646 | -99.9729
24.7683 | -99.9917
24.772 | -100.0153
24.7757 | -100.0438
24.7794 | -100.077
24.7831 | -100.1155
24.7868 | -100.1598
24.7904 | -100.2104
24.7941 | -100.2658
24.7978 | -100.3255
24.8015 | -100.39
24.8052 | -100.4589
24.8089 | -100.531
24.8126 | -100.6067
24.8163 | -100.6863
24.82 | -100.7687
24.8236 | -100.8545
24.8273 | -100.9419
24.831 | -101.0318
24.8347 | -101.1236
24.8384 | -101.2172
24.8421 | -101.3118
24.8458 | -101.4057
24.8495 | -101.5011
24.8532 | -101.5977
24.8568 | -101.6944
24.8605 | -101.7911
24.8642 | -101.8865
24.8679 | -101.9825
24.8716 | -102.0767
24.8753 | -102.1686
24.879 | -102.2581
24.8827 | -102.3445
24.8863 | -102.4292
24.89 | -102.5099
24.8937 | -102.5881
24.8974 | -102.6643
24.9011 | -102.7343
24.9048 | -102.8031
24.9085 | -102.8691
24.9122 | -102.9321
24.9159 | -102.9951
24.9195 | -103.0546
24.9232 | -103.1129
24.9269 | -103.1724
24.9306 | -103.2431
24.9343 | -103.3218
24.938 | -103.4001
24.9417 | -103.4399
24.9454 | -103.46
24.9491 | -103.4763
24.9527 | -103.489
24.9564 | -103.4972
24.9601 | -103.5042
24.9638 | -103.5122
24.9675 | -103.5191
24.9712 | -103.5241
24.9749 | -103.5262
24.9786 | -103.529
24.9822 | -103.5324
24.9859 | -103.5364
24.9896 | -103.5383
24.9933 | -103.5421
24.997 | -103.5444
25.0007 | -103.5454
25.0044 | -103.5464
25.0081 | -103.5507
25.0118 | -103.5524
25.0154 | -103.555
25.0191 | -103.5599
25.0228 | -103.5645
25.0265 | -103.566
25.0302 | -103.5693
25.0339 | -103.5748
25.0376 | -103.5798
25.0413 | -103.586
25.045 | -103.5924
25.0486 | -103.5986
25.0523 | -103.6067
25.056 | -103.614
25.0597 | -103.6208
25.0634 | -103.629
25.0671 | -103.6373
25.0708 | -103.6469
25.0745 | -103.6573
25.0782 | -103.6681
25.0818 | -103.6785
25.0855 | -103.6896
25.0892 | -103.7017
25.0929 | -103.7142
25.0966 | -103.7293
25.1003 | -103.7463
25.104 | -103.7621
25.1077 | -103.7805
25.1113 | -103.801
25.115 | -103.82
25.1187 | -103.8397
25.1224 | -103.8598
25.1261 | -103.8796
25.1298 | -103.8996
25.1335 | -103.9224
25.1372 | -103.9445
25.1409 | -103.9674
25.1445 | -103.9889
25.1482 | -104.01
25.1519 | -104.0319
25.1556 | -104.0573
25.1593 | -104.0827
25.163 | -104.1061
25.1667 | -104.1285
25.1704 | -104.1526
25.1741 | -104.1775
25.1777 | -104.2037
25.1814 | -104.2307
25.1851 | -104.2553
25.1888 | -104.2811
25.1925 | -104.3061
25.1962 | -104.329
25.1999 | -104.3548
25.2036 | -104.379
25.2072 | -104.4023
25.2109 | -104.4272
25.2146 | -104.4496
25.2183 | -104.4697
25.222 | -104.4896
25.2257 | -104.5112
25.2294 | -104.5315
25.2331 | -104.5502
25.2368 | -104.5704
25.2404 | -104.5882
25.2441 | -104.6055
25.2478 | -104.6255
25.2515 | -104.6482
25.2552 | -104.6756
25.2589 | -104.6988
25.2626 | -104.715
25.2663 | -104.7317
25.27 | -104.7481
25.2736 | -104.7612
25.2773 | -104.7757
25.281 | -104.7908
25.2847 | -104.8057
25.2884 | -104.822
25.2921 | -104.8377
25.2958 | -104.8527
25.2995 | -104.8668
25.3031 | -104.882
25.3068 | -104.8959
25.3105 | -104.9086
25.3142 | -104.924
25.3179 | -104.9365
25.3216 | -104.9456
25.3253 | -104.9583
25.329 | -104.9715
25.3327 | -104.9841
25.3363 | -104.9962
25.34 | -105.0071
25.3437 | -105.0156
25.3474 | -105.0238
25.3511 | -105.0329
25.3548 | -105.0421
25.3585 | -105.053
25.3622 | -105.0638
25.3659 | -105.0719
25.3695 | -105.0822
25.3732 | -105.0929
25.3769 | -105.1023
25.3806 | -105.1099
25.3843 | -105.1212
25.388 | -105.1319
25.3917 | -105.1406
25.3954 | -105.1472
25.399 | -105.1536
25.4027 | -105.1591
25.4064 | -105.1643
25.4101 | -105.167
25.4138 | -105.1707
25.4175 | -105.1758
25.4212 | -105.1792
25.4249 | -105.1807
25.4286 | -105.1837
25.4322 | -105.1861
25.4359 | -105.1877
25.4396 | -105.1884
25.4433 | -105.1903
25.447 | -105.19
25.4507 | -105.1899
25.4544 | -105.1891
25.4581 | -105.1855
25.4618 | -105.1819
25.4654 | -105.1802
25.4691 | -105.1774
25.4728 | -105.1756
25.4765 | -105.1752
25.4802 | -105.173
25.4839 | -105.1717
25.4876 | -105.1704
25.4913 | -105.1669
25.495 | -105.1645
25.4986 | -105.1609
25.5023 | -105.1572
25.506 | -105.1546
25.5097 | -105.1526
25.5134 | -105.1494
25.5171 | -105.1459
25.5208 | -105.1418
25.5245 | -105.1389
25.5281 | -105.1362
25.5318 | -105.1336
25.5355 | -105.1319
25.5392 | -105.136
25.5429 | -105.1347
25.5466 | -105.1352
25.5503 | -105.1357
25.554 | -105.1362
25.5577 | -105.1382
25.5613 | -105.1392
25.565 | -105.1408
25.5687 | -105.1424
25.5724 | -105.1423
25.5761 | -105.1445
25.5798 | -105.1482
25.5835 | -105.1515
25.5872 | -105.1545
25.5909 | -105.1577
25.5945 | -105.1647
25.5982 | -105.1718
25.6019 | -105.1795
25.6056 | -105.1863
25.6093 | -105.1902
25.613 | -105.1927
25.6167 | -105.1966
25.6204 | -105.202
25.624 | -105.2082
25.6277 | -105.2136
25.6314 | -105.2154
25.6351 | -105.2165
25.6388 | -105.2204
25.6425 | -105.2265
25.6462 | -105.2322
25.6499 | -105.2369
25.6536 | -105.2416
25.6572 | -105.2449
25.6609 | -105.2471
25.6646 | -105.2518
25.6683 | -105.2557
25.672 | -105.2596
25.6757 | -105.2637
25.6794 | -105.2675
25.6831 | -105.2719
25.6868 | -105.2731
25.6904 | -105.2742
25.6941 | -105.2758
25.6978 | -105.2756
25.7015 | -105.2756
25.7052 | -105.2772
25.7089 | -105.2766
25.7126 | -105.2748
25.7163 | -105.273
25.72 | -105.269
25.7236 | -105.2649
25.7273 | -105.2625
25.731 | -105.2603
25.7347 | -105.2555
25.7384 | -105.2485
25.7421 | -105.2428
25.7458 | -105.2368
25.7495 | -105.2292
25.7531 | -105.2235
25.7568 | -105.217
25.7605 | -105.2112
25.7642 | -105.2077
25.7679 | -105.2025
25.7716 | -105.1974
25.7753 | -105.1923
25.779 | -105.1874
25.7827 | -105.1809
25.7863 | -105.1751
25.79 | -105.1687
25.7937 | -105.1635
25.7974 | -105.1576
25.8011 | -105.1516
25.8048 | -105.1447
25.8085 | -105.139
25.8122 | -105.1329
25.8159 | -105.1273
25.8195 | -105.122
25.8232 | -105.1145
25.8269 | -105.1052
25.8306 | -105.0977
25.8343 | -105.0884
25.838 | -105.0788
25.8417 | -105.0723
25.8454 | -105.067
25.849 | -105.0623
25.8527 | -105.0581
25.8564 | -105.0534
25.8601 | -105.0466
25.8638 | -105.0419
25.8675 | -105.0354
25.8712 | -105.0279
25.8749 | -105.0214
25.8786 | -105.0144
25.8822 | -105.0101
25.8859 | -105.0061
25.8896 | -105.0008
25.8933 | -104.9944
25.897 | -104.989
25.9007 | -104.9825
25.9044 | -104.9746
25.9081 | -104.968
25.9118 | -104.9624
25.9154 | -104.955
25.9191 | -104.9484
25.9228 | -104.943
25.9265 | -104.9378
25.9302 | -104.9325
25.9339 | -104.9262
25.9376 | -104.9204
25.9413 | -104.9116
25.9449 | -104.9042
25.9486 | -104.8985
25.9523 | -104.8911
25.956 | -104.8838
25.9597 | -104.8765
25.9634 | -104.8675
25.9671 | -104.8591
25.9708 | -104.8511
25.9745 | -104.8431
25.9781 | -104.8365
25.9818 | -104.8293
25.9855 | -104.8225
25.9892 | -104.8158
25.9929 | -104.8073
25.9966 | -104.8005
26.0003 | -104.7936
26.004 | -104.786
26.0077 | -104.7777
26.0113 | -104.7702
26.015 | -104.7612
26.0187 | -104.7502
26.0224 | -104.7418
26.0261 | -104.7359
26.0298 | -104.7289
26.0335 | -104.7219
26.0372 | -104.7141
26.0408 | -104.7055
26.0445 | -104.6977
26.0482 | -104.6909
26.0519 | -104.6845
26.0556 | -104.679
26.0593 | -104.6733
26.063 | -104.668
26.0667 | -104.6636
26.0704 | -104.6582
26.074 | -104.6552
26.0777 | -104.6503
26.0814 | -104.6438
26.0851 | -104.6376
26.0888 | -104.6325
26.0925 | -104.6269
26.0962 | -104.6211
26.0999 | -104.6156
26.1036 | -104.6099
26.1072 | -104.6067
26.1109 | -104.6041
26.1146 | -104.5994
26.1183 | -104.5951
26.122 | -104.5905
26.1257 | -104.5869
26.1294 | -104.5829
26.1331 | -104.578
26.1368 | -104.575
26.1404 | -104.572
26.1441 | -104.5677
26.1478 | -104.5628
26.1515 | -104.5594
26.1552 | -104.5558
26.1589 | -104.5512
26.1626 | -104.5463
26.1663 | -104.5427
26.1699 | -104.5388
26.1736 | -104.5333
26.1773 | -104.5287
26.181 | -104.5257
26.1847 | -104.5208
26.1884 | -104.518
26.1921 | -104.517
26.1958 | -104.515
26.1995 | -104.5111
26.2031 | -104.5071
26.2068 | -104.5017
26.2105 | -104.4969
26.2142 | -104.4919
26.2179 | -104.4876
26.2216 | -104.4841
26.2253 | -104.4791
26.229 | -104.4723
26.2327 | -104.4664
26.2363 | -104.4636
26.24 | -104.4594
26.2437 | -104.4538
26.2474 | -104.4506
26.2511 | -104.4464
26.2548 | -104.4406
26.2585 | -104.4345
26.2622 | -104.4276
26.2658 | -104.4209
26.2695 | -104.4149
26.2732 | -104.4091
26.2769 | -104.4026
26.2806 | -104.3965
26.2843 | -104.3914
26.288 | -104.3828
26.2917 | -104.3745
26.2954 | -104.3647
26.299 | -104.355
26.3027 | -104.3485
26.3064 | -104.3413
26.3101 | -104.3346
26.3138 | -104.3265
26.3175 | -104.3182
26.3212 | -104.3092
26.3249 | -104.3009
26.3286 | -104.2911
26.3322 | -104.2799
26.3359 | -104.2703
26.3396 | -104.26
26.3433 | -104.2472
26.347 | -104.2333
26.3507 | -104.22
26.3544 | -104.2071
26.3581 | -104.1945
26.3618 | -104.183
26.3654 | -104.1689
26.3691 | -104.1556
26.3728 | -104.1424
26.3765 | -104.1286
26.3802 | -104.1141
26.3839 | -104.0998
26.3876 | -104.088
26.3913 | -104.0762
26.3949 | -104.0618
26.3986 | -104.0427
26.4023 | -104.0242
26.406 | -104.0079
26.4097 | -103.9931
26.4134 | -103.9784
26.4171 | -103.9638
26.4208 | -103.9474
26.4245 | -103.93
26.4281 | -103.9128
26.4318 | -103.8973
26.4355 | -103.8814
26.4392 | -103.8655
26.4429 | -103.85
26.4466 | -103.8345
26.4503 | -103.8189
26.454 | -103.8038
26.4577 | -103.7869
26.4613 | -103.7698
26.465 | -103.7519
26.4687 | -103.7357
26.4724 | -103.7197
26.4761 | -103.7053
26.4798 | -103.6886
26.4835 | -103.672
26.4872 | -103.6576
26.4908 | -103.6427
26.4945 | -103.6287
26.4982 | -103.6149
26.5019 | -103.6013
26.5056 | -103.5886
26.5093 | -103.5758
26.513 | -103.5623
26.5167 | -103.549
26.5204 | -103.5374
26.524 | -103.5273
26.5277 | -103.5186
26.5314 | -103.5112
26.5351 | -103.5037
26.5388 | -103.4973
26.5425 | -103.4929
26.5462 | -103.4902
26.5499 | -103.4875
26.5536 | -103.4843
26.5572 | -103.4811
26.5609 | -103.4771
26.5646 | -103.4749
26.5683 | -103.4711
26.572 | -103.4695
26.5757 | -103.4718
26.5794 | -103.4741
26.5831 | -103.4746
26.5867 | -103.4761
26.5904 | -103.4761
26.5941 | -103.4772
26.5978 | -103.4805
26.6015 | -103.4863
26.6052 | -103.4936
26.6089 | -103.4959
26.6126 | -103.4951
26.6163 | -103.4985
26.6199 | -103.503
26.6236 | -103.5063
26.6273 | -103.5097
26.631 | -103.5135
26.6347 | -103.5192
26.6384 | -103.5266
26.6421 | -103.5317
26.6458 | -103.534
26.6495 | -103.5378
26.6531 | -103.5412
26.6568 | -103.5425
26.6605 | -103.5458
26.6642 | -103.5517
26.6679 | -103.5556
26.6716 | -103.5602
26.6753 | -103.5641
26.679 | -103.5652
26.6826 | -103.5667
26.6863 | -103.5676
26.69 | -103.5652
26.6937 | -103.5637
26.6974 | -103.564
26.7011 | -103.5616
26.7048 | -103.5591
26.7085 | -103.5583
26.7122 | -103.5567
26.7158 | -103.5543
26.7195 | -103.553
26.7232 | -103.5512
26.7269 | -103.5474
26.7306 | -103.5422
26.7343 | -103.5316
26.738 | -103.521
26.7417 | -103.516
26.7454 | -103.5088
26.749 | -103.4996
26.7527 | -103.4913
26.7564 | -103.4832
26.7601 | -103.4757
26.7638 | -103.4648
26.7675 | -103.4507
26.7712 | -103.4389
26.7749 | -103.4306
26.7786 | -103.4189
26.7822 | -103.4065
26.7859 | -103.394
26.7896 | -103.38
26.7933 | -103.366
26.797 | -103.3548
26.8007 | -103.3448
26.8044 | -103.3332
26.8081 | -103.3205
26.8117 | -103.3073
26.8154 | -103.2939
26.8191 | -103.2818
26.8228 | -103.2676
26.8265 | -103.2527
26.8302 | -103.2386
26.8339 | -103.227
26.8376 | -103.2154
26.8413 | -103.2056
26.8449 | -103.1952
26.8486 | -103.1853
26.8523 | -103.1763
26.856 | -103.1679
26.8597 | -103.1564
26.8634 | -103.1461
26.8671 | -103.1377
26.8708 | -103.1284
26.8745 | -103.1193
26.8781 | -103.1109
26.8818 | -103.1044
26.8855 | -103.099
26.8892 | -103.0947
26.8929 | -103.0906
26.8966 | -103.0863
26.9003 | -103.0832
26.904 | -103.0785
26.9076 | -103.075
26.9113 | -103.0747
26.915 | -103.0719
26.9187 | -103.0692
26.9224 | -103.0677
26.9261 | -103.0656
26.9298 | -103.067
26.9335 | -103.0705
26.9372 | -103.0713
26.9408 | -103.072
26.9445 | -103.0756
26.9482 | -103.0786
26.9519 | -103.0792
26.9556 | -103.0817
26.9593 | -103.0842
26.963 | -103.0867
26.9667 | -103.0931
26.9704 | -103.0987
26.974 | -103.1046
26.9777 | -103.1116
26.9814 | -103.1167
26.9851 | -103.1201
26.9888 | -103.1271
26.9925 | -103.1356
26.9962 | -103.1407
26.9999 | -103.1452
27.0036 | -103.1488
27.0072 | -103.1519
27.0109 | -103.1555
27.0146 | -103.1576
27.0183 | -103.1571
27.022 | -103.16
27.0257 | -103.1653
27.0294 | -103.1682
27.0331 | -103.1698
27.0367 | -103.1715
27.0404 | -103.1706
27.0441 | -103.1696
27.0478 | -103.1693
27.0515 | -103.1682
27.0552 | -103.1665
27.0589 | -103.1656
27.0626 | -103.1627
27.0663 | -103.1613
27.0699 | -103.1605
27.0736 | -103.1532
27.0773 | -103.1484
27.081 | -103.1453
27.0847 | -103.1389
27.0884 | -103.1313
27.0921 | -103.1254
27.0958 | -103.1205
27.0995 | -103.1162
27.1031 | -103.1112
27.1068 | -103.1051
27.1105 | -103.0993
27.1142 | -103.0974
27.1179 | -103.0932
27.1216 | -103.0831
27.1253 | -103.0771
27.129 | -103.0717
27.1326 | -103.0648
27.1363 | -103.0618
27.14 | -103.0564
27.1437 | -103.0517
27.1474 | -103.0482
27.1511 | -103.0456
27.1548 | -103.0436
27.1585 | -103.0407
27.1622 | -103.0404
27.1658 | -103.0392
27.1695 | -103.0375
27.1732 | -103.0366
27.1769 | -103.0363
27.1806 | -103.0364
27.1843 | -103.0371
27.188 | -103.0403
27.1917 | -103.044
27.1954 | -103.0484
27.199 | -103.0539
27.2027 | -103.0584
27.2064 | -103.0655
27.2101 | -103.0739
27.2138 | -103.0842
27.2175 | -103.0938
27.2212 | -103.1024
27.2249 | -103.1109
27.2285 | -103.1178
27.2322 | -103.1255
27.2359 | -103.1345
27.2396 | -103.1451
27.2433 | -103.1585
27.247 | -103.1702
27.2507 | -103.1809
27.2544 | -103.1931
27.2581 | -103.2077
27.2617 | -103.2224
27.2654 | -103.2363
27.2691 | -103.2502
27.2728 | -103.2637
27.2765 | -103.2783
27.2802 | -103.2942
27.2839 | -103.3089
27.2876 | -103.3265
27.2913 | -103.3438
27.2949 | -103.3599
27.2986 | -103.3783
27.3023 | -103.3961
27.306 | -103.4133
27.3097 | -103.4303
27.3134 | -103.4445
27.3171 | -103.4599
27.3208 | -103.4786
27.3244 | -103.4978
27.3281 | -103.5164
27.3318 | -103.5341
27.3355 | -103.5534
27.3392 | -103.5726
27.3429 | -103.592
27.3466 | -103.6098
27.3503 | -103.6253
27.354 | -103.6418
27.3576 | -103.6571
27.3613 | -103.6719
27.365 | -103.6881
27.3687 | -103.7045
27.3724 | -103.7196
27.3761 | -103.7344
27.3798 | -103.7484
27.3835 | -103.7595
27.3872 | -103.7699
27.3908 | -103.7831
27.3945 | -103.7966
27.3982 | -103.8093
27.4019 | -103.8223
27.4056 | -103.8318
27.4093 | -103.8394
27.413 | -103.8477
27.4167 | -103.8572
27.4204 | -103.8639
27.424 | -103.8693
27.4277 | -103.8754
27.4314 | -103.8808
27.4351 | -103.8862
27.4388 | -103.8906
27.4425 | -103.8946
27.4462 | -103.8977
27.4499 | -103.8994
27.4535 | -103.9024
27.4572 | -103.9101
27.4609 | -103.9118
27.4646 | -103.9116
27.4683 | -103.9132
27.472 | -103.9142
27.4757 | -103.9123
27.4794 | -103.9101
27.4831 | -103.9068
27.4867 | -103.9038
27.4904 | -103.9009
27.4941 | -103.8976
27.4978 | -103.8931
27.5015 | -103.8899
27.5052 | -103.8897
27.5089 | -103.8856
27.5126 | -103.88
27.5163 | -103.8757
27.5199 | -103.8715
27.5236 | -103.8689
27.5273 | -103.8652
27.531 | -103.8614
27.5347 | -103.8593
27.5384 | -103.8578
27.5421 | -103.8536
27.5458 | -103.8514
27.5494 | -103.8496
27.5531 | -103.8479
27.5568 | -103.8472
27.5605 | -103.8468
27.5642 | -103.8461
27.5679 | -103.8444
27.5716 | -103.8418
27.5753 | -103.8408
27.579 | -103.841
27.5826 | -103.8421
27.5863 | -103.8447
27.59 | -103.8484
27.5937 | -103.8519
27.5974 | -103.8567
27.6011 | -103.8588
27.6048 | -103.8617
27.6085 | -103.8652
27.6122 | -103.8685
27.6158 | -103.8738
27.6195 | -103.8801
27.6232 | -103.8859
27.6269 | -103.8904
27.6306 | -103.8962
27.6343 | -103.9042
27.638 | -103.9136
27.6417 | -103.9218
27.6454 | -103.9286
27.649 | -103.9351
27.6527 | -103.9405
27.6564 | -103.946
27.6601 | -103.9523
27.6638 | -103.9598
27.6675 | -103.9604
27.6712 | -103.9648
27.6749 | -103.9682
27.6785 | -103.9708
27.6822 | -103.9745
27.6859 | -103.978
27.6896 | -103.9814
27.6933 | -103.983
27.697 | -103.9844
27.7007 | -103.9872
27.7044 | -103.9901
27.7081 | -103.9899
27.7117 | -103.9865
27.7154 | -103.9823
27.7191 | -103.9791
27.7228 | -103.9754
27.7265 | -103.9688
27.7302 | -103.9637
27.7339 | -103.9572
27.7376 | -103.9512
27.7413 | -103.9451
27.7449 | -103.9382
27.7486 | -103.9313
27.7523 | -103.9244
27.756 | -103.9173
27.7597 | -103.9112
27.7634 | -103.9041
27.7671 | -103.8967
27.7708 | -103.8883
27.7744 | -103.8819
27.7781 | -103.8766
27.7818 | -103.867
27.7855 | -103.8555
27.7892 | -103.8438
27.7929 | -103.8347
27.7966 | -103.8255
27.8003 | -103.817
27.804 | -103.809
27.8076 | -103.8006
27.8113 | -103.7924
27.815 | -103.7839
27.8187 | -103.7739
27.8224 | -103.7658
27.8261 | -103.7578
27.8298 | -103.7503
27.8335 | -103.7428
27.8372 | -103.737
27.8408 | -103.7322
27.8445 | -103.7279
27.8482 | -103.7239
27.8519 | -103.7206
27.8556 | -103.7149
27.8593 | -103.7052
27.863 | -103.6973
27.8667 | -103.6936
27.8703 | -103.6933
27.874 | -103.6926
27.8777 | -103.6913
27.8814 | -103.6924
27.8851 | -103.6945
27.8888 | -103.6968
27.8925 | -103.6976
27.8962 | -103.698
27.8999 | -103.6993
27.9035 | -103.7017
27.9072 | -103.7038
27.9109 | -103.7059
27.9146 | -103.7066
27.9183 | -103.7073
27.922 | -103.7056
27.9257 | -103.695
27.9294 | -103.6689
27.9331 | -103.6535
27.9367 | -103.6451
27.9404 | -103.6365
27.9441 | -103.6286
27.9478 | -103.6206
27.9515 | -103.6131
27.9552 | -103.605
27.9589 | -103.5968
27.9626 | -103.588
27.9662 | -103.5767
27.9699 | -103.5648
27.9736 | -103.551
27.9773 | -103.5353
27.981 | -103.5189
27.9847 | -103.4994
27.9884 | -103.4801
27.9921 | -103.462
27.9958 | -103.4419
27.9994 | -103.4181
28.0031 | -103.3937